## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 208 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 205
Hood | 02-92B/P | 145
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 999.3 | hPa |  | PRC PRO 01
Humidity | 42.5 | %/Hr
Internal temperature of calibration bench | 23.13 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1007 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.3 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.87 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.01 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3129 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 13/04/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 19/04/2018
Engineer | Anthony HUE | Approval |  | 23/04/2018
CEO | Robert GLEMOT | Approval |  | 23/04/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4324 | 57.4224
0.012209 | -36.0088 | 38.2166
0.018313 | -34.9589 | 28.0344
0.024417 | -34.5022 | 21.7923
0.030521 | -34.3918 | 18.0228
0.036626 | -34.2484 | 14.9743
0.04273 | -34.1539 | 12.6212
0.048834 | -34.1093 | 11.1189
0.054938 | -34.0462 | 9.9829
0.061043 | -33.9854 | 8.8945
0.067147 | -33.9775 | 8.056
0.073251 | -33.9809 | 7.2808
0.079355 | -34.0012 | 6.6103
0.08546 | -34.0165 | 6.0155
0.091564 | -34.0129 | 5.5784
0.097668 | -33.9958 | 5.3296
0.10377 | -33.9754 | 5.0728
0.10988 | -33.9598 | 4.7896
0.11598 | -33.9519 | 4.4933
0.12209 | -33.9545 | 4.2089
0.12819 | -33.9651 | 3.9259
0.13429 | -33.9764 | 3.6399
0.1404 | -33.9794 | 3.3668
0.1465 | -33.9811 | 3.1137
0.15261 | -33.9784 | 2.8968
0.15871 | -33.9726 | 2.7352
0.16482 | -33.9657 | 2.6297
0.17092 | -33.9651 | 2.5863
0.17702 | -33.9643 | 2.5113
0.18313 | -33.9644 | 2.4313
0.18923 | -33.9646 | 2.3213
0.19534 | -33.9638 | 2.1938
0.20144 | -33.9627 | 2.0502
0.20754 | -33.9602 | 1.9289
0.21365 | -33.9595 | 1.8144
0.21975 | -33.9594 | 1.7176
0.22586 | -33.961 | 1.6324
0.23196 | -33.9639 | 1.5631
0.23807 | -33.9683 | 1.5034
0.24417 | -33.9727 | 1.445
0.25027 | -33.9771 | 1.3826
0.25638 | -33.9798 | 1.3196
0.26248 | -33.9821 | 1.2511
0.26859 | -33.9836 | 1.1826
0.27469 | -33.9846 | 1.1082
0.2808 | -33.9862 | 1.032
0.2869 | -33.9876 | 0.95906
0.293 | -33.989 | 0.8901
0.29911 | -33.9905 | 0.82915
0.30521 | -33.992 | 0.7698
0.31132 | -33.9932 | 0.73225
0.31742 | -33.995 | 0.68988
0.32353 | -33.9959 | 0.65552
0.32963 | -33.9974 | 0.62678
0.33573 | -33.9989 | 0.60653
0.34184 | -34.0008 | 0.58467
0.34794 | -34.0034 | 0.55923
0.35405 | -34.0061 | 0.53377
0.36015 | -34.0089 | 0.50694
0.36626 | -34.0113 | 0.47767
0.37236 | -34.0138 | 0.43934
0.37846 | -34.0163 | 0.39351
0.38457 | -34.0173 | 0.33595
0.39067 | -34.0185 | 0.28831
0.39678 | -34.0193 | 0.23603
0.40288 | -34.0198 | 0.18166
0.40899 | -34.0197 | 0.12956
0.41509 | -34.0191 | 0.078687
0.42119 | -34.0191 | 0.035333
0.4273 | -34.019 | -0.011259
0.4334 | -34.019 | -0.057735
0.43951 | -34.0193 | -0.10177
0.44561 | -34.0202 | -0.14231
0.45172 | -34.022 | -0.17545
0.45782 | -34.0229 | -0.21344
0.46392 | -34.0243 | -0.25088
0.47003 | -34.0255 | -0.28443
0.47613 | -34.0266 | -0.31966
0.48224 | -34.0274 | -0.35564
0.48834 | -34.0268 | -0.39165
0.49445 | -34.0268 | -0.42796
0.50055 | -34.0266 | -0.46107
0.50665 | -34.0261 | -0.49632
0.51276 | -34.0254 | -0.53134
0.51886 | -34.0247 | -0.56595
0.52497 | -34.0237 | -0.59069
0.53107 | -34.0232 | -0.62092
0.53717 | -34.0228 | -0.64981
0.54328 | -34.0222 | -0.67345
0.54938 | -34.0222 | -0.6965
0.55549 | -34.0224 | -0.71801
0.56159 | -34.0238 | -0.73038
0.5677 | -34.0243 | -0.74567
0.5738 | -34.0256 | -0.7626
0.5799 | -34.0266 | -0.77834
0.58601 | -34.028 | -0.78931
0.59211 | -34.0299 | -0.79651
0.59822 | -34.0314 | -0.80682
0.60432 | -34.0332 | -0.81985
0.61043 | -34.0347 | -0.8287
0.61653 | -34.0362 | -0.83976
0.62263 | -34.038 | -0.84995
0.62874 | -34.0388 | -0.86135
0.63484 | -34.0399 | -0.87132
0.64095 | -34.0407 | -0.88214
0.64705 | -34.0413 | -0.89266
0.65316 | -34.0417 | -0.90581
0.65926 | -34.042 | -0.92225
0.66536 | -34.0412 | -0.94023
0.67147 | -34.041 | -0.95418
0.67757 | -34.0409 | -0.96918
0.68368 | -34.0404 | -0.98864
0.68978 | -34.0399 | -1.0064
0.69589 | -34.0393 | -1.0258
0.70199 | -34.0386 | -1.0487
0.70809 | -34.038 | -1.0681
0.7142 | -34.0376 | -1.0888
0.7203 | -34.0368 | -1.1084
0.72641 | -34.0364 | -1.131
0.73251 | -34.036 | -1.1512
0.73862 | -34.0364 | -1.173
0.74472 | -34.0363 | -1.1931
0.75082 | -34.0363 | -1.2152
0.75693 | -34.0368 | -1.2371
0.76303 | -34.0376 | -1.2591
0.76914 | -34.0395 | -1.2796
0.77524 | -34.0407 | -1.2994
0.78135 | -34.0421 | -1.3195
0.78745 | -34.0436 | -1.3396
0.79355 | -34.045 | -1.3596
0.79966 | -34.0469 | -1.3792
0.80576 | -34.0492 | -1.3994
0.81187 | -34.0509 | -1.4213
0.81797 | -34.053 | -1.4429
0.82408 | -34.0547 | -1.4638
0.83018 | -34.0566 | -1.4847
0.83628 | -34.0586 | -1.5037
0.84239 | -34.0599 | -1.525
0.84849 | -34.0614 | -1.5473
0.8546 | -34.063 | -1.5677
0.8607 | -34.0643 | -1.589
0.86681 | -34.0655 | -1.6086
0.87291 | -34.0666 | -1.6282
0.87901 | -34.0672 | -1.647
0.88512 | -34.068 | -1.6659
0.89122 | -34.0689 | -1.6849
0.89733 | -34.0694 | -1.7014
0.90343 | -34.07 | -1.7196
0.90953 | -34.0703 | -1.7347
0.91564 | -34.0705 | -1.7514
0.92174 | -34.0709 | -1.7671
0.92785 | -34.0711 | -1.7824
0.93395 | -34.0714 | -1.7969
0.94006 | -34.0717 | -1.8104
0.94616 | -34.0717 | -1.8228
0.95226 | -34.072 | -1.8361
0.95837 | -34.0722 | -1.8505
0.96447 | -34.0726 | -1.8639
0.97058 | -34.0726 | -1.8777
0.97668 | -34.0726 | -1.8915
0.98279 | -34.0725 | -1.9047
0.98889 | -34.0725 | -1.9159
0.99499 | -34.0726 | -1.9298
1.0011 | -34.0727 | -1.941
1.0072 | -34.0725 | -1.9531
1.0133 | -34.0724 | -1.9643
1.0194 | -34.0721 | -1.9778
1.0255 | -34.072 | -1.9919
1.0316 | -34.0716 | -2.0043
1.0377 | -34.0714 | -2.0183
1.0438 | -34.071 | -2.0315
1.0499 | -34.0704 | -2.0471
1.056 | -34.0699 | -2.0618
1.0621 | -34.0694 | -2.0765
1.0682 | -34.0691 | -2.0894
1.0743 | -34.0686 | -2.1053
1.0805 | -34.0683 | -2.1201
1.0866 | -34.0675 | -2.1356
1.0927 | -34.0668 | -2.1505
1.0988 | -34.0663 | -2.1658
1.1049 | -34.0656 | -2.1803
1.111 | -34.065 | -2.1961
1.1171 | -34.0644 | -2.211
1.1232 | -34.0641 | -2.2266
1.1293 | -34.0638 | -2.241
1.1354 | -34.0636 | -2.2564
1.1415 | -34.0634 | -2.2724
1.1476 | -34.0629 | -2.2883
1.1537 | -34.0625 | -2.3034
1.1598 | -34.0627 | -2.3185
1.1659 | -34.0626 | -2.3323
1.172 | -34.0624 | -2.3466
1.1781 | -34.0623 | -2.361
1.1842 | -34.0622 | -2.3756
1.1903 | -34.0627 | -2.3871
1.1964 | -34.0627 | -2.4009
1.2025 | -34.0629 | -2.4153
1.2086 | -34.0633 | -2.4281
1.2147 | -34.0635 | -2.441
1.2209 | -34.0639 | -2.4527
1.227 | -34.0647 | -2.4638
1.2331 | -34.0653 | -2.4763
1.2392 | -34.0657 | -2.4886
1.2453 | -34.0663 | -2.5001
1.2514 | -34.0668 | -2.5105
1.2575 | -34.0675 | -2.5218
1.2636 | -34.0684 | -2.5313
1.2697 | -34.0691 | -2.541
1.2758 | -34.0699 | -2.5526
1.2819 | -34.0707 | -2.563
1.288 | -34.0715 | -2.5742
1.2941 | -34.0724 | -2.5832
1.3002 | -34.0732 | -2.5917
1.3063 | -34.0739 | -2.6012
1.3124 | -34.0746 | -2.6107
1.3185 | -34.0755 | -2.6202
1.3246 | -34.0763 | -2.6298
1.3307 | -34.0772 | -2.6389
1.3368 | -34.0781 | -2.6481
1.3429 | -34.0789 | -2.6576
1.349 | -34.0795 | -2.6673
1.3551 | -34.0803 | -2.6771
1.3613 | -34.081 | -2.6859
1.3674 | -34.0817 | -2.6964
1.3735 | -34.0822 | -2.7067
1.3796 | -34.0828 | -2.7162
1.3857 | -34.0832 | -2.726
1.3918 | -34.0839 | -2.7366
1.3979 | -34.0845 | -2.7462
1.404 | -34.0851 | -2.7581
1.4101 | -34.0856 | -2.7689
1.4162 | -34.0861 | -2.7799
1.4223 | -34.0867 | -2.7918
1.4284 | -34.0873 | -2.8022
1.4345 | -34.0876 | -2.8142
1.4406 | -34.088 | -2.828
1.4467 | -34.0886 | -2.8402
1.4528 | -34.0891 | -2.852
1.4589 | -34.0897 | -2.8643
1.465 | -34.09 | -2.8756
1.4711 | -34.0906 | -2.8874
1.4772 | -34.0912 | -2.9011
1.4833 | -34.0917 | -2.9145
1.4894 | -34.0923 | -2.9272
1.4955 | -34.093 | -2.9395
1.5016 | -34.0936 | -2.9524
1.5078 | -34.094 | -2.9667
1.5139 | -34.0945 | -2.9801
1.52 | -34.0951 | -2.9931
1.5261 | -34.0955 | -3.006
1.5322 | -34.0959 | -3.0204
1.5383 | -34.0963 | -3.0334
1.5444 | -34.097 | -3.0463
1.5505 | -34.0975 | -3.0601
1.5566 | -34.0981 | -3.0735
1.5627 | -34.0987 | -3.0876
1.5688 | -34.0994 | -3.1016
1.5749 | -34.0998 | -3.1145
1.581 | -34.1002 | -3.1282
1.5871 | -34.1007 | -3.1423
1.5932 | -34.1012 | -3.1555
1.5993 | -34.1017 | -3.17
1.6054 | -34.1021 | -3.1836
1.6115 | -34.1026 | -3.1981
1.6176 | -34.103 | -3.211
1.6237 | -34.1033 | -3.2249
1.6298 | -34.1037 | -3.2378
1.6359 | -34.1042 | -3.2508
1.642 | -34.1047 | -3.2631
1.6482 | -34.105 | -3.2764
1.6543 | -34.1054 | -3.2903
1.6604 | -34.1058 | -3.3027
1.6665 | -34.106 | -3.3155
1.6726 | -34.1063 | -3.3283
1.6787 | -34.1068 | -3.3412
1.6848 | -34.1068 | -3.354
1.6909 | -34.1071 | -3.3671
1.697 | -34.1075 | -3.3796
1.7031 | -34.1076 | -3.3925
1.7092 | -34.1077 | -3.4049
1.7153 | -34.1079 | -3.4172
1.7214 | -34.1078 | -3.4298
1.7275 | -34.1079 | -3.4426
1.7336 | -34.1079 | -3.4559
1.7397 | -34.1082 | -3.4685
1.7458 | -34.1082 | -3.481
1.7519 | -34.1084 | -3.4935
1.758 | -34.1085 | -3.5057
1.7641 | -34.1087 | -3.5177
1.7702 | -34.1088 | -3.5309
1.7763 | -34.1088 | -3.5427
1.7824 | -34.1089 | -3.5554
1.7885 | -34.109 | -3.5682
1.7947 | -34.1092 | -3.5805
1.8008 | -34.1093 | -3.5921
1.8069 | -34.1093 | -3.604
1.813 | -34.1095 | -3.6165
1.8191 | -34.1096 | -3.6295
1.8252 | -34.1097 | -3.6426
1.8313 | -34.11 | -3.6552
1.8374 | -34.1102 | -3.6674
1.8435 | -34.1103 | -3.6796
1.8496 | -34.1105 | -3.693
1.8557 | -34.1107 | -3.7061
1.8618 | -34.1111 | -3.7187
1.8679 | -34.1113 | -3.7311
1.874 | -34.1115 | -3.7435
1.8801 | -34.1117 | -3.7563
1.8862 | -34.1121 | -3.7686
1.8923 | -34.1123 | -3.7812
1.8984 | -34.1125 | -3.7936
1.9045 | -34.1126 | -3.8056
1.9106 | -34.1128 | -3.8182
1.9167 | -34.1129 | -3.8308
1.9228 | -34.1131 | -3.843
1.9289 | -34.1133 | -3.8554
1.9351 | -34.1136 | -3.8667
1.9412 | -34.1139 | -3.8784
1.9473 | -34.1142 | -3.8909
1.9534 | -34.1145 | -3.9027
1.9595 | -34.1148 | -3.9149
1.9656 | -34.1151 | -3.9277
1.9717 | -34.1154 | -3.9398
1.9778 | -34.1158 | -3.9517
1.9839 | -34.1162 | -3.9636
1.99 | -34.1164 | -3.9752
1.9961 | -34.1166 | -3.9877
2.0022 | -34.1169 | -3.9997
2.0083 | -34.1171 | -4.0113
2.0144 | -34.1174 | -4.0231
2.0205 | -34.1178 | -4.0354
2.0266 | -34.1182 | -4.0472
2.0327 | -34.1185 | -4.0593
2.0388 | -34.119 | -4.0713
2.0449 | -34.1193 | -4.0832
2.051 | -34.1196 | -4.095
2.0571 | -34.1201 | -4.1068
2.0632 | -34.1203 | -4.1188
2.0693 | -34.1206 | -4.1314
2.0754 | -34.121 | -4.1432
2.0816 | -34.1214 | -4.1558
2.0877 | -34.1216 | -4.1678
2.0938 | -34.1221 | -4.1803
2.0999 | -34.1223 | -4.1925
2.106 | -34.1226 | -4.2046
2.1121 | -34.1229 | -4.2169
2.1182 | -34.1232 | -4.2282
2.1243 | -34.1236 | -4.2394
2.1304 | -34.1239 | -4.251
2.1365 | -34.1243 | -4.2633
2.1426 | -34.1246 | -4.2757
2.1487 | -34.125 | -4.2873
2.1548 | -34.1255 | -4.2996
2.1609 | -34.126 | -4.3115
2.167 | -34.1264 | -4.3231
2.1731 | -34.1268 | -4.3356
2.1792 | -34.1271 | -4.3471
2.1853 | -34.1276 | -4.3595
2.1914 | -34.128 | -4.3713
2.1975 | -34.1284 | -4.3833
2.2036 | -34.1287 | -4.3958
2.2097 | -34.1291 | -4.4077
2.2158 | -34.1295 | -4.4188
2.222 | -34.1299 | -4.4297
2.2281 | -34.1302 | -4.4415
2.2342 | -34.1306 | -4.4533
2.2403 | -34.1309 | -4.4644
2.2464 | -34.1312 | -4.4755
2.2525 | -34.1316 | -4.4873
2.2586 | -34.132 | -4.4983
2.2647 | -34.1324 | -4.5092
2.2708 | -34.1328 | -4.5208
2.2769 | -34.1332 | -4.5319
2.283 | -34.1334 | -4.5435
2.2891 | -34.1338 | -4.5548
2.2952 | -34.134 | -4.5657
2.3013 | -34.1343 | -4.577
2.3074 | -34.1347 | -4.5878
2.3135 | -34.135 | -4.5996
2.3196 | -34.1353 | -4.6102
2.3257 | -34.1355 | -4.6213
2.3318 | -34.1358 | -4.6328
2.3379 | -34.136 | -4.644
2.344 | -34.1362 | -4.6558
2.3501 | -34.1365 | -4.6669
2.3562 | -34.1368 | -4.6776
2.3623 | -34.137 | -4.6884
2.3685 | -34.1372 | -4.6993
2.3746 | -34.1374 | -4.7109
2.3807 | -34.1377 | -4.7218
2.3868 | -34.1379 | -4.7329
2.3929 | -34.1382 | -4.7439
2.399 | -34.1384 | -4.7548
2.4051 | -34.1387 | -4.7657
2.4112 | -34.139 | -4.7767
2.4173 | -34.1392 | -4.7877
2.4234 | -34.1394 | -4.7989
2.4295 | -34.1396 | -4.8101
2.4356 | -34.1398 | -4.8213
2.4417 | -34.14 | -4.8329
2.4478 | -34.1402 | -4.8441
2.4539 | -34.1405 | -4.8547
2.46 | -34.1406 | -4.8665
2.4661 | -34.1409 | -4.8779
2.4722 | -34.1411 | -4.8885
2.4783 | -34.1413 | -4.8993
2.4844 | -34.1414 | -4.9106
2.4905 | -34.1416 | -4.922
2.4966 | -34.1418 | -4.9331
2.5027 | -34.142 | -4.944
2.5089 | -34.1422 | -4.9552
2.515 | -34.1423 | -4.9664
2.5211 | -34.1424 | -4.9774
2.5272 | -34.1425 | -4.9887
2.5333 | -34.1427 | -5
2.5394 | -34.143 | -5.0111
2.5455 | -34.1432 | -5.0228
2.5516 | -34.1434 | -5.0335
2.5577 | -34.1436 | -5.0453
2.5638 | -34.1438 | -5.0573
2.5699 | -34.1439 | -5.0691
2.576 | -34.1441 | -5.081
2.5821 | -34.1442 | -5.0925
2.5882 | -34.1444 | -5.104
2.5943 | -34.1446 | -5.1152
2.6004 | -34.1448 | -5.1265
2.6065 | -34.145 | -5.1372
2.6126 | -34.1452 | -5.148
2.6187 | -34.1454 | -5.1588
2.6248 | -34.1456 | -5.1699
2.6309 | -34.1458 | -5.1817
2.637 | -34.1461 | -5.1936
2.6431 | -34.1463 | -5.2044
2.6492 | -34.1465 | -5.2161
2.6554 | -34.1468 | -5.2273
2.6615 | -34.1469 | -5.2386
2.6676 | -34.1472 | -5.2504
2.6737 | -34.1475 | -5.2619
2.6798 | -34.1479 | -5.2733
2.6859 | -34.1481 | -5.2849
2.692 | -34.1483 | -5.2961
2.6981 | -34.1486 | -5.3076
2.7042 | -34.1489 | -5.319
2.7103 | -34.1491 | -5.33
2.7164 | -34.1495 | -5.3413
2.7225 | -34.1498 | -5.353
2.7286 | -34.1501 | -5.3643
2.7347 | -34.1503 | -5.3752
2.7408 | -34.1507 | -5.3857
2.7469 | -34.1509 | -5.3973
2.753 | -34.1513 | -5.4082
2.7591 | -34.1516 | -5.4198
2.7652 | -34.1519 | -5.4316
2.7713 | -34.1521 | -5.4427
2.7774 | -34.1525 | -5.454
2.7835 | -34.1527 | -5.4653
2.7896 | -34.153 | -5.4762
2.7958 | -34.1533 | -5.4873
2.8019 | -34.1536 | -5.498
2.808 | -34.154 | -5.5091
2.8141 | -34.1543 | -5.5206
2.8202 | -34.1547 | -5.5318
2.8263 | -34.1551 | -5.5431
2.8324 | -34.1555 | -5.5547
2.8385 | -34.1559 | -5.5659
2.8446 | -34.1563 | -5.577
2.8507 | -34.1568 | -5.5882
2.8568 | -34.1572 | -5.599
2.8629 | -34.1575 | -5.61
2.869 | -34.1579 | -5.6214
2.8751 | -34.1583 | -5.6325
2.8812 | -34.1588 | -5.6442
2.8873 | -34.1591 | -5.6554
2.8934 | -34.1595 | -5.6668
2.8995 | -34.1599 | -5.6782
2.9056 | -34.1602 | -5.6897
2.9117 | -34.1606 | -5.7011
2.9178 | -34.1611 | -5.7121
2.9239 | -34.1615 | -5.724
2.93 | -34.162 | -5.7358
2.9361 | -34.1624 | -5.747
2.9423 | -34.1628 | -5.758
2.9484 | -34.1632 | -5.7694
2.9545 | -34.1637 | -5.7807
2.9606 | -34.1642 | -5.7925
2.9667 | -34.1646 | -5.804
2.9728 | -34.165 | -5.8156
2.9789 | -34.1655 | -5.827
2.985 | -34.166 | -5.8387
2.9911 | -34.1664 | -5.8504
2.9972 | -34.1669 | -5.8618
3.0033 | -34.1674 | -5.8737
3.0094 | -34.1678 | -5.8851
3.0155 | -34.1683 | -5.8964
3.0216 | -34.1688 | -5.908
3.0277 | -34.1692 | -5.92
3.0338 | -34.1696 | -5.9317
3.0399 | -34.1701 | -5.9434
3.046 | -34.1705 | -5.9553
3.0521 | -34.1709 | -5.9668
3.0582 | -34.1712 | -5.9792
3.0643 | -34.1717 | -5.991
3.0704 | -34.172 | -6.0032
3.0765 | -34.1724 | -6.0147
3.0827 | -34.1728 | -6.0263
3.0888 | -34.1732 | -6.0379
3.0949 | -34.1736 | -6.0496
3.101 | -34.1741 | -6.0615
3.1071 | -34.1744 | -6.0736
3.1132 | -34.1749 | -6.0857
3.1193 | -34.1753 | -6.0975
3.1254 | -34.1758 | -6.1092
3.1315 | -34.1762 | -6.1215
3.1376 | -34.1765 | -6.1333
3.1437 | -34.177 | -6.1452
3.1498 | -34.1774 | -6.1576
3.1559 | -34.1778 | -6.1694
3.162 | -34.1781 | -6.1819
3.1681 | -34.1784 | -6.1942
3.1742 | -34.1788 | -6.2061
3.1803 | -34.1792 | -6.2181
3.1864 | -34.1796 | -6.2298
3.1925 | -34.18 | -6.2417
3.1986 | -34.1803 | -6.2541
3.2047 | -34.1807 | -6.2665
3.2108 | -34.181 | -6.278
3.2169 | -34.1814 | -6.2901
3.223 | -34.1817 | -6.3022
3.2292 | -34.182 | -6.3144
3.2353 | -34.1823 | -6.3256
3.2414 | -34.1826 | -6.3374
3.2475 | -34.1829 | -6.3492
3.2536 | -34.1832 | -6.361
3.2597 | -34.1836 | -6.3724
3.2658 | -34.1838 | -6.3837
3.2719 | -34.1841 | -6.3951
3.278 | -34.1844 | -6.4071
3.2841 | -34.1848 | -6.4185
3.2902 | -34.1851 | -6.4304
3.2963 | -34.1854 | -6.4422
3.3024 | -34.1856 | -6.4537
3.3085 | -34.1859 | -6.4656
3.3146 | -34.1861 | -6.4772
3.3207 | -34.1865 | -6.489
3.3268 | -34.1867 | -6.5009
3.3329 | -34.187 | -6.512
3.339 | -34.1872 | -6.5232
3.3451 | -34.1875 | -6.5348
3.3512 | -34.1877 | -6.5459
3.3573 | -34.188 | -6.5575
3.3634 | -34.1883 | -6.5691
3.3696 | -34.1885 | -6.5808
3.3757 | -34.1887 | -6.592
3.3818 | -34.189 | -6.6032
3.3879 | -34.1893 | -6.6147
3.394 | -34.1895 | -6.6264
3.4001 | -34.1897 | -6.6376
3.4062 | -34.1899 | -6.6489
3.4123 | -34.1901 | -6.6599
3.4184 | -34.1904 | -6.6716
3.4245 | -34.1907 | -6.6828
3.4306 | -34.1908 | -6.6939
3.4367 | -34.191 | -6.7048
3.4428 | -34.1913 | -6.7155
3.4489 | -34.1914 | -6.7269
3.455 | -34.1917 | -6.7381
3.4611 | -34.1919 | -6.7497
3.4672 | -34.1921 | -6.7611
3.4733 | -34.1923 | -6.7726
3.4794 | -34.1926 | -6.7837
3.4855 | -34.1928 | -6.7945
3.4916 | -34.193 | -6.8053
3.4977 | -34.1932 | -6.817
3.5038 | -34.1934 | -6.8282
3.5099 | -34.1936 | -6.8389
3.5161 | -34.1938 | -6.8496
3.5222 | -34.194 | -6.8604
3.5283 | -34.1942 | -6.8715
3.5344 | -34.1944 | -6.8827
3.5405 | -34.1946 | -6.8936
3.5466 | -34.1948 | -6.9049
3.5527 | -34.195 | -6.9161
3.5588 | -34.1951 | -6.9271
3.5649 | -34.1953 | -6.9381
3.571 | -34.1955 | -6.9491
3.5771 | -34.1956 | -6.9605
3.5832 | -34.1958 | -6.9708
3.5893 | -34.1961 | -6.9818
3.5954 | -34.1963 | -6.993
3.6015 | -34.1966 | -7.0042
3.6076 | -34.1968 | -7.0152
3.6137 | -34.197 | -7.0257
3.6198 | -34.1974 | -7.0363
3.6259 | -34.1975 | -7.0475
3.632 | -34.1978 | -7.0585
3.6381 | -34.198 | -7.0697
3.6442 | -34.1982 | -7.0806
3.6503 | -34.1985 | -7.0911
3.6565 | -34.1987 | -7.1022
3.6626 | -34.1989 | -7.1132
3.6687 | -34.1992 | -7.1243
3.6748 | -34.1995 | -7.1351
3.6809 | -34.1998 | -7.1461
3.687 | -34.2 | -7.1574
3.6931 | -34.2004 | -7.168
3.6992 | -34.2006 | -7.179
3.7053 | -34.2009 | -7.1898
3.7114 | -34.201 | -7.2009
3.7175 | -34.2013 | -7.2119
3.7236 | -34.2016 | -7.2223
3.7297 | -34.2019 | -7.2334
3.7358 | -34.2022 | -7.2447
3.7419 | -34.2024 | -7.2559
3.748 | -34.2027 | -7.2668
3.7541 | -34.203 | -7.2778
3.7602 | -34.2033 | -7.2886
3.7663 | -34.2035 | -7.2994
3.7724 | -34.2039 | -7.3104
3.7785 | -34.2042 | -7.3212
3.7846 | -34.2045 | -7.332
3.7907 | -34.2048 | -7.3432
3.7969 | -34.2051 | -7.3543
3.803 | -34.2054 | -7.3649
3.8091 | -34.2057 | -7.3758
3.8152 | -34.2061 | -7.3866
3.8213 | -34.2063 | -7.3978
3.8274 | -34.2067 | -7.4085
3.8335 | -34.2071 | -7.4195
3.8396 | -34.2074 | -7.4305
3.8457 | -34.2078 | -7.4416
3.8518 | -34.2082 | -7.4525
3.8579 | -34.2086 | -7.4634
3.864 | -34.2089 | -7.4745
3.8701 | -34.2093 | -7.4859
3.8762 | -34.2096 | -7.4966
3.8823 | -34.2099 | -7.5078
3.8884 | -34.2103 | -7.5188
3.8945 | -34.2107 | -7.5301
3.9006 | -34.211 | -7.5414
3.9067 | -34.2114 | -7.5528
3.9128 | -34.2117 | -7.5639
3.9189 | -34.2121 | -7.5747
3.925 | -34.2124 | -7.5859
3.9311 | -34.2127 | -7.5964
3.9372 | -34.2131 | -7.6077
3.9434 | -34.2134 | -7.6187
3.9495 | -34.2139 | -7.6299
3.9556 | -34.2143 | -7.6407
3.9617 | -34.2146 | -7.6514
3.9678 | -34.215 | -7.6621
3.9739 | -34.2153 | -7.6731
3.98 | -34.2157 | -7.6843
3.9861 | -34.216 | -7.6952
3.9922 | -34.2163 | -7.7062
3.9983 | -34.2167 | -7.7168
4.0044 | -34.2171 | -7.7274
4.0105 | -34.2174 | -7.7384
4.0166 | -34.2178 | -7.7495
4.0227 | -34.2182 | -7.7604
4.0288 | -34.2185 | -7.7719
4.0349 | -34.2188 | -7.7827
4.041 | -34.2191 | -7.7936
4.0471 | -34.2194 | -7.8047
4.0532 | -34.2198 | -7.8153
4.0593 | -34.2202 | -7.8265
4.0654 | -34.2205 | -7.8372
4.0715 | -34.2208 | -7.8487
4.0776 | -34.2212 | -7.8596
4.0838 | -34.2216 | -7.8709
4.0899 | -34.222 | -7.8817
4.096 | -34.2223 | -7.8926
4.1021 | -34.2226 | -7.9033
4.1082 | -34.2229 | -7.9141
4.1143 | -34.2233 | -7.9247
4.1204 | -34.2236 | -7.9354
4.1265 | -34.224 | -7.946
4.1326 | -34.2243 | -7.9566
4.1387 | -34.2247 | -7.9673
4.1448 | -34.225 | -7.9783
4.1509 | -34.2253 | -7.9892
4.157 | -34.2256 | -8.0005
4.1631 | -34.226 | -8.0117
4.1692 | -34.2263 | -8.0224
4.1753 | -34.2267 | -8.0332
4.1814 | -34.2269 | -8.0437
4.1875 | -34.2273 | -8.0549
4.1936 | -34.2276 | -8.0659
4.1997 | -34.2279 | -8.0764
4.2058 | -34.2282 | -8.0875
4.2119 | -34.2285 | -8.0982
4.218 | -34.2288 | -8.1088
4.2241 | -34.2291 | -8.1194
4.2303 | -34.2295 | -8.1301
4.2364 | -34.2298 | -8.1409
4.2425 | -34.2301 | -8.1517
4.2486 | -34.2303 | -8.1627
4.2547 | -34.2305 | -8.1733
4.2608 | -34.2308 | -8.1842
4.2669 | -34.2311 | -8.1947
4.273 | -34.2314 | -8.2056
4.2791 | -34.2317 | -8.2164
4.2852 | -34.232 | -8.2274
4.2913 | -34.2322 | -8.2379
4.2974 | -34.2325 | -8.2489
4.3035 | -34.2328 | -8.26
4.3096 | -34.2331 | -8.2709
4.3157 | -34.2334 | -8.2815
4.3218 | -34.2337 | -8.2922
4.3279 | -34.234 | -8.303
4.334 | -34.2342 | -8.3136
4.3401 | -34.2345 | -8.324
4.3462 | -34.2348 | -8.3348
4.3523 | -34.2351 | -8.3451
4.3584 | -34.2353 | -8.3554
4.3645 | -34.2356 | -8.3661
4.3707 | -34.2359 | -8.3765
4.3768 | -34.2362 | -8.3871
4.3829 | -34.2365 | -8.3982
4.389 | -34.2367 | -8.4088
4.3951 | -34.237 | -8.4193
4.4012 | -34.2372 | -8.4303
4.4073 | -34.2375 | -8.4408
4.4134 | -34.2378 | -8.4514
4.4195 | -34.2381 | -8.462
4.4256 | -34.2384 | -8.4724
4.4317 | -34.2387 | -8.483
4.4378 | -34.2389 | -8.4933
4.4439 | -34.2392 | -8.5037
4.45 | -34.2394 | -8.5143
4.4561 | -34.2396 | -8.5249
4.4622 | -34.2399 | -8.5356
4.4683 | -34.2401 | -8.5461
4.4744 | -34.2404 | -8.557
4.4805 | -34.2407 | -8.5673
4.4866 | -34.2409 | -8.5778
4.4927 | -34.2412 | -8.5887
4.4988 | -34.2415 | -8.5992
4.5049 | -34.2418 | -8.6094
4.511 | -34.2421 | -8.6197
4.5172 | -34.2423 | -8.6302
4.5233 | -34.2425 | -8.6412
4.5294 | -34.2428 | -8.6516
4.5355 | -34.2431 | -8.6622
4.5416 | -34.2434 | -8.6727
4.5477 | -34.2436 | -8.6833
4.5538 | -34.2439 | -8.6934
4.5599 | -34.2442 | -8.7039
4.566 | -34.2444 | -8.7145
4.5721 | -34.2447 | -8.725
4.5782 | -34.245 | -8.7358
4.5843 | -34.2453 | -8.7465
4.5904 | -34.2456 | -8.7571
4.5965 | -34.2459 | -8.7675
4.6026 | -34.2461 | -8.7785
4.6087 | -34.2464 | -8.7889
4.6148 | -34.2466 | -8.7998
4.6209 | -34.2469 | -8.8103
4.627 | -34.2472 | -8.8211
4.6331 | -34.2474 | -8.8315
4.6392 | -34.2477 | -8.8422
4.6453 | -34.248 | -8.8527
4.6514 | -34.2483 | -8.8633
4.6576 | -34.2485 | -8.8743
4.6637 | -34.2488 | -8.8845
4.6698 | -34.249 | -8.8949
4.6759 | -34.2493 | -8.9059
4.682 | -34.2495 | -8.9163
4.6881 | -34.2497 | -8.9269
4.6942 | -34.25 | -8.9374
4.7003 | -34.2503 | -8.948
4.7064 | -34.2506 | -8.9584
4.7125 | -34.2509 | -8.9692
4.7186 | -34.2511 | -8.9795
4.7247 | -34.2514 | -8.9902
4.7308 | -34.2517 | -9.0007
4.7369 | -34.252 | -9.0114
4.743 | -34.2522 | -9.0223
4.7491 | -34.2525 | -9.0332
4.7552 | -34.2528 | -9.0439
4.7613 | -34.253 | -9.0546
4.7674 | -34.2532 | -9.0651
4.7735 | -34.2535 | -9.0761
4.7796 | -34.2538 | -9.087
4.7857 | -34.2541 | -9.0977
4.7918 | -34.2544 | -9.1082
4.7979 | -34.2547 | -9.1187
4.8041 | -34.255 | -9.1294
4.8102 | -34.2553 | -9.1397
4.8163 | -34.2555 | -9.1502
4.8224 | -34.2558 | -9.1609
4.8285 | -34.2561 | -9.1712
4.8346 | -34.2564 | -9.1816
4.8407 | -34.2567 | -9.1923
4.8468 | -34.257 | -9.2032
4.8529 | -34.2573 | -9.2141
4.859 | -34.2576 | -9.225
4.8651 | -34.2578 | -9.2354
4.8712 | -34.2581 | -9.2458
4.8773 | -34.2584 | -9.2566
4.8834 | -34.2587 | -9.2672
4.8895 | -34.259 | -9.2776
4.8956 | -34.2592 | -9.2883
4.9017 | -34.2594 | -9.2987
4.9078 | -34.2597 | -9.3096
4.9139 | -34.26 | -9.3199
4.92 | -34.2604 | -9.3302
4.9261 | -34.2607 | -9.3405
4.9322 | -34.261 | -9.3511
4.9383 | -34.2613 | -9.3617
4.9445 | -34.2616 | -9.3724
4.9506 | -34.2618 | -9.3828
4.9567 | -34.2621 | -9.3932
4.9628 | -34.2624 | -9.4037
4.9689 | -34.2627 | -9.4147
4.975 | -34.263 | -9.4256
4.9811 | -34.2633 | -9.436
4.9872 | -34.2636 | -9.4465
4.9933 | -34.264 | -9.4573
4.9994 | -34.2643 | -9.4683
5.0055 | -34.2646 | -9.479
5.0116 | -34.2649 | -9.4901
5.0177 | -34.2653 | -9.501
5.0238 | -34.2656 | -9.5119
5.0299 | -34.266 | -9.5229
5.036 | -34.2663 | -9.5336
5.0421 | -34.2666 | -9.5441
5.0482 | -34.2669 | -9.5548
5.0543 | -34.2672 | -9.5652
5.0604 | -34.2676 | -9.5759
5.0665 | -34.2679 | -9.5863
5.0726 | -34.2682 | -9.5972
5.0787 | -34.2685 | -9.6081
5.0848 | -34.2689 | -9.6189
5.091 | -34.2692 | -9.6294
5.0971 | -34.2695 | -9.6397
5.1032 | -34.2698 | -9.6507
5.1093 | -34.2702 | -9.6614
5.1154 | -34.2705 | -9.6718
5.1215 | -34.2707 | -9.6823
5.1276 | -34.2711 | -9.6927
5.1337 | -34.2715 | -9.7031
5.1398 | -34.2718 | -9.7135
5.1459 | -34.2721 | -9.7244
5.152 | -34.2724 | -9.7352
5.1581 | -34.2727 | -9.7461
5.1642 | -34.2731 | -9.7568
5.1703 | -34.2734 | -9.7675
5.1764 | -34.2738 | -9.778
5.1825 | -34.2741 | -9.7884
5.1886 | -34.2744 | -9.7988
5.1947 | -34.2747 | -9.8095
5.2008 | -34.2751 | -9.8203
5.2069 | -34.2754 | -9.8312
5.213 | -34.2758 | -9.8417
5.2191 | -34.2761 | -9.8524
5.2252 | -34.2764 | -9.8631
5.2314 | -34.2767 | -9.874
5.2375 | -34.2771 | -9.8844
5.2436 | -34.2774 | -9.8951
5.2497 | -34.2777 | -9.9056
5.2558 | -34.2781 | -9.9166
5.2619 | -34.2784 | -9.9275
5.268 | -34.2788 | -9.9384
5.2741 | -34.2792 | -9.9491
5.2802 | -34.2795 | -9.9602
5.2863 | -34.2799 | -9.9708
5.2924 | -34.2802 | -9.9817
5.2985 | -34.2806 | -9.9925
5.3046 | -34.2809 | -10.0033
5.3107 | -34.2813 | -10.0142
5.3168 | -34.2816 | -10.025
5.3229 | -34.282 | -10.0357
5.329 | -34.2823 | -10.0465
5.3351 | -34.2827 | -10.0569
5.3412 | -34.283 | -10.0676
5.3473 | -34.2833 | -10.0784
5.3534 | -34.2837 | -10.0893
5.3595 | -34.284 | -10.1001
5.3656 | -34.2844 | -10.1108
5.3717 | -34.2848 | -10.1219
5.3779 | -34.2851 | -10.133
5.384 | -34.2855 | -10.1437
5.3901 | -34.2858 | -10.1547
5.3962 | -34.2862 | -10.1653
5.4023 | -34.2865 | -10.1763
5.4084 | -34.2869 | -10.187
5.4145 | -34.2873 | -10.1975
5.4206 | -34.2876 | -10.2083
5.4267 | -34.288 | -10.2194
5.4328 | -34.2884 | -10.2303
5.4389 | -34.2887 | -10.2408
5.445 | -34.289 | -10.252
5.4511 | -34.2894 | -10.2625
5.4572 | -34.2898 | -10.2734
5.4633 | -34.2901 | -10.2842
5.4694 | -34.2904 | -10.295
5.4755 | -34.2908 | -10.3055
5.4816 | -34.2912 | -10.3163
5.4877 | -34.2916 | -10.327
5.4938 | -34.2919 | -10.3375
5.4999 | -34.2923 | -10.3482
5.506 | -34.2927 | -10.359
5.5121 | -34.293 | -10.3696
5.5183 | -34.2933 | -10.3803
5.5244 | -34.2937 | -10.3913
5.5305 | -34.2941 | -10.402
5.5366 | -34.2944 | -10.4125
5.5427 | -34.2947 | -10.4234
5.5488 | -34.2951 | -10.4342
5.5549 | -34.2954 | -10.4451
5.561 | -34.2958 | -10.4561
5.5671 | -34.2962 | -10.4667
5.5732 | -34.2965 | -10.4777
5.5793 | -34.2969 | -10.4889
5.5854 | -34.2973 | -10.4996
5.5915 | -34.2976 | -10.5107
5.5976 | -34.298 | -10.5216
5.6037 | -34.2983 | -10.5324
5.6098 | -34.2986 | -10.5432
5.6159 | -34.299 | -10.5538
5.622 | -34.2994 | -10.5647
5.6281 | -34.2997 | -10.5756
5.6342 | -34.3001 | -10.5867
5.6403 | -34.3005 | -10.5975
5.6464 | -34.3008 | -10.6085
5.6525 | -34.3012 | -10.6194
5.6586 | -34.3016 | -10.6299
5.6648 | -34.302 | -10.6407
5.6709 | -34.3023 | -10.6517
5.677 | -34.3027 | -10.6625
5.6831 | -34.3031 | -10.6734
5.6892 | -34.3034 | -10.6845
5.6953 | -34.3038 | -10.6953
5.7014 | -34.3041 | -10.7062
5.7075 | -34.3044 | -10.717
5.7136 | -34.3048 | -10.7278
5.7197 | -34.3051 | -10.7386
5.7258 | -34.3055 | -10.7492
5.7319 | -34.3059 | -10.7601
5.738 | -34.3062 | -10.7709
5.7441 | -34.3066 | -10.7816
5.7502 | -34.307 | -10.7921
5.7563 | -34.3074 | -10.803
5.7624 | -34.3078 | -10.814
5.7685 | -34.3081 | -10.8248
5.7746 | -34.3085 | -10.8359
5.7807 | -34.3088 | -10.8468
5.7868 | -34.3091 | -10.8578
5.7929 | -34.3095 | -10.8685
5.799 | -34.3098 | -10.8792
5.8052 | -34.3102 | -10.8899
5.8113 | -34.3105 | -10.9006
5.8174 | -34.3109 | -10.9113
5.8235 | -34.3112 | -10.922
5.8296 | -34.3116 | -10.933
5.8357 | -34.3119 | -10.9438
5.8418 | -34.3123 | -10.9547
5.8479 | -34.3126 | -10.9656
5.854 | -34.313 | -10.9765
5.8601 | -34.3133 | -10.9872
5.8662 | -34.3137 | -10.9981
5.8723 | -34.314 | -11.0087
5.8784 | -34.3143 | -11.0195
5.8845 | -34.3147 | -11.0302
5.8906 | -34.315 | -11.041
5.8967 | -34.3154 | -11.0516
5.9028 | -34.3158 | -11.0622
5.9089 | -34.3162 | -11.0729
5.915 | -34.3165 | -11.0834
5.9211 | -34.3169 | -11.0942
5.9272 | -34.3172 | -11.1047
5.9333 | -34.3176 | -11.1154
5.9394 | -34.3179 | -11.1261
5.9455 | -34.3183 | -11.1367
5.9517 | -34.3187 | -11.1472
5.9578 | -34.3191 | -11.1581
5.9639 | -34.3195 | -11.1685
5.97 | -34.3199 | -11.1793
5.9761 | -34.3203 | -11.1899
5.9822 | -34.3206 | -11.2007
5.9883 | -34.321 | -11.2115
5.9944 | -34.3213 | -11.2223
6.0005 | -34.3217 | -11.2332
6.0066 | -34.3221 | -11.2437
6.0127 | -34.3225 | -11.2541
6.0188 | -34.3228 | -11.2651
6.0249 | -34.3232 | -11.2759
6.031 | -34.3235 | -11.2865
6.0371 | -34.3239 | -11.2972
6.0432 | -34.3242 | -11.3078
6.0493 | -34.3246 | -11.3183
6.0554 | -34.325 | -11.3288
6.0615 | -34.3254 | -11.3395
6.0676 | -34.3257 | -11.3504
6.0737 | -34.3261 | -11.3612
6.0798 | -34.3265 | -11.3719
6.0859 | -34.3269 | -11.3825
6.0921 | -34.3272 | -11.3929
6.0982 | -34.3276 | -11.4036
6.1043 | -34.328 | -11.4141
6.1104 | -34.3284 | -11.425
6.1165 | -34.3288 | -11.4352
6.1226 | -34.3291 | -11.4458
6.1287 | -34.3295 | -11.4564
6.1348 | -34.3299 | -11.467
6.1409 | -34.3302 | -11.4776
6.147 | -34.3306 | -11.4882
6.1531 | -34.3309 | -11.4987
6.1592 | -34.3313 | -11.5093
6.1653 | -34.3317 | -11.5198
6.1714 | -34.332 | -11.5304
6.1775 | -34.3324 | -11.5411
6.1836 | -34.3328 | -11.5518
6.1897 | -34.3332 | -11.5624
6.1958 | -34.3335 | -11.5732
6.2019 | -34.3339 | -11.584
6.208 | -34.3343 | -11.5946
6.2141 | -34.3347 | -11.6051
6.2202 | -34.3351 | -11.6159
6.2263 | -34.3355 | -11.6264
6.2325 | -34.3358 | -11.6366
6.2386 | -34.3362 | -11.6469
6.2447 | -34.3366 | -11.6578
6.2508 | -34.3369 | -11.6684
6.2569 | -34.3374 | -11.6789
6.263 | -34.3377 | -11.6895
6.2691 | -34.3381 | -11.7
6.2752 | -34.3384 | -11.7106
6.2813 | -34.3388 | -11.7212
6.2874 | -34.3391 | -11.7316
6.2935 | -34.3395 | -11.7421
6.2996 | -34.3398 | -11.7527
6.3057 | -34.3402 | -11.7633
6.3118 | -34.3405 | -11.774
6.3179 | -34.3409 | -11.7845
6.324 | -34.3413 | -11.795
6.3301 | -34.3416 | -11.8055
6.3362 | -34.342 | -11.8159
6.3423 | -34.3424 | -11.8266
6.3484 | -34.3428 | -11.8373
6.3545 | -34.3432 | -11.8475
6.3606 | -34.3435 | -11.8577
6.3667 | -34.3438 | -11.8681
6.3728 | -34.3442 | -11.8787
6.379 | -34.3446 | -11.8892
6.3851 | -34.345 | -11.8995
6.3912 | -34.3453 | -11.91
6.3973 | -34.3457 | -11.9205
6.4034 | -34.346 | -11.9308
6.4095 | -34.3464 | -11.9413
6.4156 | -34.3467 | -11.9518
6.4217 | -34.3471 | -11.9623
6.4278 | -34.3475 | -11.9729
6.4339 | -34.3479 | -11.9833
6.44 | -34.3483 | -11.9937
6.4461 | -34.3486 | -12.0044
6.4522 | -34.349 | -12.0146
6.4583 | -34.3493 | -12.0253
6.4644 | -34.3496 | -12.0359
6.4705 | -34.35 | -12.046
6.4766 | -34.3504 | -12.0565
6.4827 | -34.3508 | -12.067
6.4888 | -34.3511 | -12.0773
6.4949 | -34.3515 | -12.0875
6.501 | -34.3518 | -12.098
6.5071 | -34.3522 | -12.1082
6.5132 | -34.3526 | -12.1187
6.5194 | -34.3529 | -12.1292
6.5255 | -34.3533 | -12.1394
6.5316 | -34.3536 | -12.15
6.5377 | -34.354 | -12.1604
6.5438 | -34.3543 | -12.1706
6.5499 | -34.3547 | -12.181
6.556 | -34.3551 | -12.1914
6.5621 | -34.3554 | -12.202
6.5682 | -34.3558 | -12.2125
6.5743 | -34.3562 | -12.2228
6.5804 | -34.3565 | -12.2334
6.5865 | -34.3569 | -12.2438
6.5926 | -34.3573 | -12.2541
6.5987 | -34.3576 | -12.2642
6.6048 | -34.358 | -12.2747
6.6109 | -34.3584 | -12.285
6.617 | -34.3587 | -12.2956
6.6231 | -34.3591 | -12.3059
6.6292 | -34.3595 | -12.3163
6.6353 | -34.3598 | -12.327
6.6414 | -34.3602 | -12.3373
6.6475 | -34.3606 | -12.3477
6.6536 | -34.3609 | -12.3579
6.6597 | -34.3613 | -12.368
6.6659 | -34.3616 | -12.3785
6.672 | -34.362 | -12.3887
6.6781 | -34.3624 | -12.3991
6.6842 | -34.3628 | -12.4093
6.6903 | -34.3632 | -12.4197
6.6964 | -34.3635 | -12.4301
6.7025 | -34.3639 | -12.4403
6.7086 | -34.3642 | -12.4506
6.7147 | -34.3646 | -12.461
6.7208 | -34.3649 | -12.4715
6.7269 | -34.3653 | -12.4818
6.733 | -34.3656 | -12.4919
6.7391 | -34.366 | -12.5023
6.7452 | -34.3664 | -12.5124
6.7513 | -34.3668 | -12.5224
6.7574 | -34.3671 | -12.5329
6.7635 | -34.3675 | -12.543
6.7696 | -34.3679 | -12.5532
6.7757 | -34.3683 | -12.5635
6.7818 | -34.3687 | -12.5739
6.7879 | -34.3691 | -12.584
6.794 | -34.3694 | -12.5941
6.8001 | -34.3697 | -12.6043
6.8063 | -34.3701 | -12.6146
6.8124 | -34.3705 | -12.6251
6.8185 | -34.3708 | -12.6355
6.8246 | -34.3712 | -12.6456
6.8307 | -34.3715 | -12.6558
6.8368 | -34.3719 | -12.6663
6.8429 | -34.3723 | -12.6766
6.849 | -34.3727 | -12.687
6.8551 | -34.3731 | -12.6971
6.8612 | -34.3734 | -12.7074
6.8673 | -34.3738 | -12.7176
6.8734 | -34.3741 | -12.728
6.8795 | -34.3745 | -12.7383
6.8856 | -34.3749 | -12.7486
6.8917 | -34.3753 | -12.7588
6.8978 | -34.3756 | -12.7688
6.9039 | -34.376 | -12.7788
6.91 | -34.3764 | -12.7891
6.9161 | -34.3767 | -12.7991
6.9222 | -34.3771 | -12.8093
6.9283 | -34.3775 | -12.8193
6.9344 | -34.3779 | -12.8296
6.9405 | -34.3783 | -12.8398
6.9466 | -34.3787 | -12.8502
6.9528 | -34.3791 | -12.8606
6.9589 | -34.3795 | -12.871
6.965 | -34.3798 | -12.8813
6.9711 | -34.3802 | -12.8916
6.9772 | -34.3807 | -12.9019
6.9833 | -34.381 | -12.9121
6.9894 | -34.3814 | -12.9225
6.9955 | -34.3818 | -12.9325
7.0016 | -34.3821 | -12.9427
7.0077 | -34.3825 | -12.9529
7.0138 | -34.3829 | -12.9631
7.0199 | -34.3833 | -12.9732
7.026 | -34.3837 | -12.9834
7.0321 | -34.3841 | -12.9934
7.0382 | -34.3844 | -13.0032
7.0443 | -34.3848 | -13.0136
7.0504 | -34.3852 | -13.0239
7.0565 | -34.3856 | -13.0342
7.0626 | -34.386 | -13.0441
7.0687 | -34.3864 | -13.0539
7.0748 | -34.3868 | -13.0638
7.0809 | -34.3872 | -13.074
7.087 | -34.3875 | -13.0842
7.0932 | -34.3879 | -13.0944
7.0993 | -34.3883 | -13.1046
7.1054 | -34.3887 | -13.1146
7.1115 | -34.3891 | -13.1245
7.1176 | -34.3894 | -13.1346
7.1237 | -34.3898 | -13.1447
7.1298 | -34.3902 | -13.1548
7.1359 | -34.3906 | -13.1649
7.142 | -34.3909 | -13.1748
7.1481 | -34.3913 | -13.1849
7.1542 | -34.3917 | -13.1951
7.1603 | -34.3921 | -13.2051
7.1664 | -34.3925 | -13.2152
7.1725 | -34.3929 | -13.2254
7.1786 | -34.3933 | -13.2355
7.1847 | -34.3937 | -13.2457
7.1908 | -34.3941 | -13.256
7.1969 | -34.3944 | -13.2663
7.203 | -34.3948 | -13.2763
7.2091 | -34.3952 | -13.2863
7.2152 | -34.3956 | -13.2964
7.2213 | -34.396 | -13.3064
7.2274 | -34.3964 | -13.3164
7.2335 | -34.3968 | -13.3265
7.2397 | -34.3972 | -13.3366
7.2458 | -34.3976 | -13.3465
7.2519 | -34.398 | -13.3566
7.258 | -34.3983 | -13.3668
7.2641 | -34.3987 | -13.3769
7.2702 | -34.3991 | -13.3872
7.2763 | -34.3995 | -13.3974
7.2824 | -34.3999 | -13.4076
7.2885 | -34.4003 | -13.4178
7.2946 | -34.4007 | -13.4279
7.3007 | -34.4011 | -13.4381
7.3068 | -34.4014 | -13.4484
7.3129 | -34.4018 | -13.4587
7.319 | -34.4023 | -13.4688
7.3251 | -34.4026 | -13.4788
7.3312 | -34.403 | -13.4889
7.3373 | -34.4034 | -13.4988
7.3434 | -34.4039 | -13.5087
7.3495 | -34.4043 | -13.5187
7.3556 | -34.4047 | -13.5289
7.3617 | -34.4051 | -13.5392
7.3678 | -34.4054 | -13.5492
7.3739 | -34.4058 | -13.5596
7.3801 | -34.4062 | -13.5699
7.3862 | -34.4066 | -13.58
7.3923 | -34.407 | -13.5902
7.3984 | -34.4074 | -13.6001
7.4045 | -34.4078 | -13.6101
7.4106 | -34.4083 | -13.6202
7.4167 | -34.4086 | -13.6304
7.4228 | -34.4091 | -13.6406
7.4289 | -34.4095 | -13.6507
7.435 | -34.4099 | -13.6606
7.4411 | -34.4103 | -13.6708
7.4472 | -34.4106 | -13.6809
7.4533 | -34.4111 | -13.6909
7.4594 | -34.4115 | -13.7012
7.4655 | -34.4119 | -13.7111
7.4716 | -34.4123 | -13.7214
7.4777 | -34.4127 | -13.7315
7.4838 | -34.4131 | -13.7415
7.4899 | -34.4135 | -13.7515
7.496 | -34.4138 | -13.7615
7.5021 | -34.4143 | -13.7715
7.5082 | -34.4147 | -13.7816
7.5143 | -34.4151 | -13.7918
7.5204 | -34.4155 | -13.802
7.5266 | -34.4159 | -13.812
7.5327 | -34.4163 | -13.822
7.5388 | -34.4166 | -13.8324
7.5449 | -34.4171 | -13.8427
7.551 | -34.4175 | -13.8527
7.5571 | -34.4179 | -13.8629
7.5632 | -34.4183 | -13.8731
7.5693 | -34.4187 | -13.8831
7.5754 | -34.4191 | -13.8931
7.5815 | -34.4195 | -13.9032
7.5876 | -34.4199 | -13.9135
7.5937 | -34.4204 | -13.9238
7.5998 | -34.4208 | -13.9339
7.6059 | -34.4212 | -13.9441
7.612 | -34.4216 | -13.9542
7.6181 | -34.422 | -13.9643
7.6242 | -34.4224 | -13.9747
7.6303 | -34.4229 | -13.985
7.6364 | -34.4232 | -13.9952
7.6425 | -34.4236 | -14.0053
7.6486 | -34.424 | -14.0155
7.6547 | -34.4244 | -14.0255
7.6608 | -34.4248 | -14.0356
7.667 | -34.4252 | -14.0459
7.6731 | -34.4256 | -14.0561
7.6792 | -34.4261 | -14.0665
7.6853 | -34.4265 | -14.0767
7.6914 | -34.4269 | -14.087
7.6975 | -34.4273 | -14.0972
7.7036 | -34.4277 | -14.1074
7.7097 | -34.4282 | -14.1177
7.7158 | -34.4286 | -14.128
7.7219 | -34.429 | -14.138
7.728 | -34.4294 | -14.1483
7.7341 | -34.4298 | -14.1586
7.7402 | -34.4302 | -14.1688
7.7463 | -34.4306 | -14.179
7.7524 | -34.431 | -14.1895
7.7585 | -34.4315 | -14.1997
7.7646 | -34.4319 | -14.2098
7.7707 | -34.4323 | -14.2202
7.7768 | -34.4327 | -14.2306
7.7829 | -34.4331 | -14.241
7.789 | -34.4335 | -14.2515
7.7951 | -34.4339 | -14.2616
7.8012 | -34.4343 | -14.2718
7.8073 | -34.4347 | -14.2823
7.8135 | -34.4352 | -14.2925
7.8196 | -34.4356 | -14.3029
7.8257 | -34.436 | -14.3133
7.8318 | -34.4364 | -14.3237
7.8379 | -34.4368 | -14.3338
7.844 | -34.4372 | -14.344
7.8501 | -34.4376 | -14.3543
7.8562 | -34.438 | -14.3648
7.8623 | -34.4385 | -14.3749
7.8684 | -34.4389 | -14.3853
7.8745 | -34.4393 | -14.3955
7.8806 | -34.4397 | -14.406
7.8867 | -34.4401 | -14.416
7.8928 | -34.4406 | -14.426
7.8989 | -34.441 | -14.4363
7.905 | -34.4414 | -14.4466
7.9111 | -34.4418 | -14.4569
7.9172 | -34.4422 | -14.4672
7.9233 | -34.4426 | -14.4775
7.9294 | -34.443 | -14.4878
7.9355 | -34.4434 | -14.4979
7.9416 | -34.4438 | -14.5082
7.9477 | -34.4442 | -14.5183
7.9539 | -34.4446 | -14.5285
7.96 | -34.4451 | -14.5389
7.9661 | -34.4455 | -14.5492
7.9722 | -34.4459 | -14.5594
7.9783 | -34.4464 | -14.5699
7.9844 | -34.4468 | -14.5801
7.9905 | -34.4472 | -14.5906
7.9966 | -34.4476 | -14.601
8.0027 | -34.448 | -14.6112
8.0088 | -34.4484 | -14.6215
8.0149 | -34.4488 | -14.6318
8.021 | -34.4492 | -14.6422
8.0271 | -34.4497 | -14.6527
8.0332 | -34.4501 | -14.663
8.0393 | -34.4505 | -14.6734
8.0454 | -34.451 | -14.6839
8.0515 | -34.4514 | -14.6942
8.0576 | -34.4518 | -14.7048
8.0637 | -34.4522 | -14.715
8.0698 | -34.4526 | -14.7252
8.0759 | -34.453 | -14.7353
8.082 | -34.4534 | -14.7455
8.0881 | -34.4538 | -14.7559
8.0942 | -34.4543 | -14.7663
8.1004 | -34.4547 | -14.7766
8.1065 | -34.4551 | -14.7868
8.1126 | -34.4555 | -14.7971
8.1187 | -34.4559 | -14.8075
8.1248 | -34.4563 | -14.8176
8.1309 | -34.4568 | -14.8278
8.137 | -34.4572 | -14.8379
8.1431 | -34.4576 | -14.8484
8.1492 | -34.458 | -14.8587
8.1553 | -34.4584 | -14.869
8.1614 | -34.4589 | -14.8795
8.1675 | -34.4593 | -14.8898
8.1736 | -34.4596 | -14.9003
8.1797 | -34.4601 | -14.9106
8.1858 | -34.4605 | -14.921
8.1919 | -34.4609 | -14.9313
8.198 | -34.4613 | -14.9414
8.2041 | -34.4618 | -14.9518
8.2102 | -34.4622 | -14.9621
8.2163 | -34.4626 | -14.9722
8.2224 | -34.463 | -14.9825
8.2285 | -34.4634 | -14.9925
8.2346 | -34.4638 | -15.0028
8.2408 | -34.4643 | -15.0129
8.2469 | -34.4647 | -15.0233
8.253 | -34.4651 | -15.0334
8.2591 | -34.4656 | -15.0438
8.2652 | -34.466 | -15.0541
8.2713 | -34.4664 | -15.0644
8.2774 | -34.4669 | -15.0748
8.2835 | -34.4673 | -15.0853
8.2896 | -34.4677 | -15.0954
8.2957 | -34.4681 | -15.1058
8.3018 | -34.4686 | -15.1161
8.3079 | -34.469 | -15.1262
8.314 | -34.4694 | -15.1365
8.3201 | -34.4699 | -15.1467
8.3262 | -34.4703 | -15.157
8.3323 | -34.4707 | -15.1673
8.3384 | -34.4711 | -15.1777
8.3445 | -34.4715 | -15.1879
8.3506 | -34.472 | -15.1982
8.3567 | -34.4724 | -15.2086
8.3628 | -34.4728 | -15.2188
8.3689 | -34.4732 | -15.2289
8.375 | -34.4737 | -15.2393
8.3812 | -34.4741 | -15.2495
8.3873 | -34.4745 | -15.2598
8.3934 | -34.4749 | -15.2702
8.3995 | -34.4753 | -15.2803
8.4056 | -34.4758 | -15.2905
8.4117 | -34.4762 | -15.3009
8.4178 | -34.4766 | -15.3112
8.4239 | -34.477 | -15.3213
8.43 | -34.4774 | -15.3316
8.4361 | -34.4778 | -15.342
8.4422 | -34.4782 | -15.3521
8.4483 | -34.4787 | -15.3622
8.4544 | -34.4791 | -15.3725
8.4605 | -34.4795 | -15.3827
8.4666 | -34.4799 | -15.3931
8.4727 | -34.4804 | -15.4034
8.4788 | -34.4808 | -15.4136
8.4849 | -34.4812 | -15.4239
8.491 | -34.4816 | -15.4342
8.4971 | -34.482 | -15.4442
8.5032 | -34.4824 | -15.4543
8.5093 | -34.4829 | -15.4644
8.5154 | -34.4833 | -15.4747
8.5215 | -34.4837 | -15.4849
8.5277 | -34.4841 | -15.4951
8.5338 | -34.4846 | -15.5053
8.5399 | -34.485 | -15.5155
8.546 | -34.4854 | -15.5255
8.5521 | -34.4859 | -15.5358
8.5582 | -34.4862 | -15.5459
8.5643 | -34.4867 | -15.5559
8.5704 | -34.4871 | -15.5659
8.5765 | -34.4875 | -15.5762
8.5826 | -34.488 | -15.5861
8.5887 | -34.4884 | -15.5962
8.5948 | -34.4888 | -15.6065
8.6009 | -34.4893 | -15.6165
8.607 | -34.4897 | -15.6268
8.6131 | -34.4901 | -15.6371
8.6192 | -34.4906 | -15.647
8.6253 | -34.491 | -15.6572
8.6314 | -34.4914 | -15.6673
8.6375 | -34.4918 | -15.6775
8.6436 | -34.4922 | -15.6877
8.6497 | -34.4926 | -15.6977
8.6558 | -34.493 | -15.7078
8.6619 | -34.4934 | -15.718
8.6681 | -34.4939 | -15.7283
8.6742 | -34.4943 | -15.7384
8.6803 | -34.4947 | -15.7485
8.6864 | -34.4952 | -15.7583
8.6925 | -34.4956 | -15.7683
8.6986 | -34.496 | -15.7785
8.7047 | -34.4964 | -15.7889
8.7108 | -34.4968 | -15.799
8.7169 | -34.4972 | -15.8092
8.723 | -34.4977 | -15.8191
8.7291 | -34.4981 | -15.8291
8.7352 | -34.4985 | -15.8394
8.7413 | -34.4989 | -15.8496
8.7474 | -34.4993 | -15.8595
8.7535 | -34.4998 | -15.8697
8.7596 | -34.5002 | -15.8798
8.7657 | -34.5007 | -15.89
8.7718 | -34.5011 | -15.9002
8.7779 | -34.5015 | -15.9104
8.784 | -34.5019 | -15.9205
8.7901 | -34.5024 | -15.9306
8.7962 | -34.5028 | -15.9405
8.8023 | -34.5032 | -15.9505
8.8084 | -34.5036 | -15.9606
8.8146 | -34.5041 | -15.9707
8.8207 | -34.5045 | -15.9807
8.8268 | -34.5049 | -15.9904
8.8329 | -34.5053 | -16.0005
8.839 | -34.5058 | -16.0106
8.8451 | -34.5062 | -16.0206
8.8512 | -34.5066 | -16.0306
8.8573 | -34.507 | -16.0406
8.8634 | -34.5075 | -16.0504
8.8695 | -34.5079 | -16.0607
8.8756 | -34.5083 | -16.0708
8.8817 | -34.5087 | -16.0807
8.8878 | -34.5091 | -16.0909
8.8939 | -34.5096 | -16.101
8.9 | -34.51 | -16.111
8.9061 | -34.5104 | -16.1211
8.9122 | -34.5109 | -16.1311
8.9183 | -34.5113 | -16.1412
8.9244 | -34.5117 | -16.1513
8.9305 | -34.5122 | -16.1611
8.9366 | -34.5126 | -16.1711
8.9427 | -34.513 | -16.1812
8.9488 | -34.5135 | -16.1912
8.955 | -34.5139 | -16.2012
8.9611 | -34.5144 | -16.2113
8.9672 | -34.5148 | -16.2212
8.9733 | -34.5153 | -16.2311
8.9794 | -34.5157 | -16.241
8.9855 | -34.5161 | -16.2512
8.9916 | -34.5166 | -16.2614
8.9977 | -34.517 | -16.2714
9.0038 | -34.5175 | -16.2812
9.0099 | -34.5179 | -16.2913
9.016 | -34.5183 | -16.3014
9.0221 | -34.5188 | -16.3112
9.0282 | -34.5192 | -16.3211
9.0343 | -34.5197 | -16.3312
9.0404 | -34.5201 | -16.3411
9.0465 | -34.5206 | -16.3513
9.0526 | -34.521 | -16.3612
9.0587 | -34.5215 | -16.3713
9.0648 | -34.5219 | -16.3812
9.0709 | -34.5223 | -16.3912
9.077 | -34.5228 | -16.4012
9.0831 | -34.5233 | -16.4114
9.0892 | -34.5237 | -16.4214
9.0953 | -34.5242 | -16.4315
9.1015 | -34.5246 | -16.4415
9.1076 | -34.5251 | -16.4513
9.1137 | -34.5255 | -16.4613
9.1198 | -34.526 | -16.4712
9.1259 | -34.5264 | -16.4812
9.132 | -34.5269 | -16.4913
9.1381 | -34.5273 | -16.5012
9.1442 | -34.5278 | -16.5108
9.1503 | -34.5282 | -16.5209
9.1564 | -34.5287 | -16.5309
9.1625 | -34.5291 | -16.5409
9.1686 | -34.5296 | -16.5508
9.1747 | -34.5301 | -16.5608
9.1808 | -34.5305 | -16.5707
9.1869 | -34.531 | -16.5806
9.193 | -34.5314 | -16.5905
9.1991 | -34.5319 | -16.6006
9.2052 | -34.5323 | -16.6106
9.2113 | -34.5328 | -16.6206
9.2174 | -34.5333 | -16.6305
9.2235 | -34.5337 | -16.6405
9.2296 | -34.5342 | -16.6504
9.2357 | -34.5346 | -16.6605
9.2419 | -34.5351 | -16.6704
9.248 | -34.5356 | -16.6802
9.2541 | -34.5361 | -16.6902
9.2602 | -34.5365 | -16.7003
9.2663 | -34.537 | -16.71
9.2724 | -34.5375 | -16.7199
9.2785 | -34.5379 | -16.7298
9.2846 | -34.5384 | -16.7396
9.2907 | -34.5388 | -16.7496
9.2968 | -34.5393 | -16.7596
9.3029 | -34.5398 | -16.7697
9.309 | -34.5402 | -16.7797
9.3151 | -34.5406 | -16.7896
9.3212 | -34.5412 | -16.7993
9.3273 | -34.5416 | -16.8091
9.3334 | -34.5421 | -16.8189
9.3395 | -34.5425 | -16.8287
9.3456 | -34.5429 | -16.8385
9.3517 | -34.5434 | -16.8485
9.3578 | -34.5439 | -16.8584
9.3639 | -34.5444 | -16.8682
9.37 | -34.5448 | -16.8781
9.3761 | -34.5453 | -16.8882
9.3822 | -34.5458 | -16.8981
9.3884 | -34.5463 | -16.9078
9.3945 | -34.5467 | -16.9177
9.4006 | -34.5472 | -16.9277
9.4067 | -34.5477 | -16.9376
9.4128 | -34.5481 | -16.9476
9.4189 | -34.5486 | -16.9577
9.425 | -34.5491 | -16.9675
9.4311 | -34.5496 | -16.9773
9.4372 | -34.55 | -16.9872
9.4433 | -34.5505 | -16.9971
9.4494 | -34.551 | -17.0068
9.4555 | -34.5515 | -17.0168
9.4616 | -34.5519 | -17.0266
9.4677 | -34.5524 | -17.0365
9.4738 | -34.5529 | -17.0464
9.4799 | -34.5533 | -17.0562
9.486 | -34.5538 | -17.0662
9.4921 | -34.5543 | -17.076
9.4982 | -34.5548 | -17.0856
9.5043 | -34.5552 | -17.0954
9.5104 | -34.5557 | -17.1052
9.5165 | -34.5562 | -17.115
9.5226 | -34.5566 | -17.125
9.5288 | -34.5571 | -17.1347
9.5349 | -34.5576 | -17.1447
9.541 | -34.5581 | -17.1543
9.5471 | -34.5585 | -17.1642
9.5532 | -34.559 | -17.1741
9.5593 | -34.5595 | -17.1841
9.5654 | -34.5599 | -17.1939
9.5715 | -34.5604 | -17.2037
9.5776 | -34.5609 | -17.2138
9.5837 | -34.5613 | -17.2236
9.5898 | -34.5618 | -17.2336
9.5959 | -34.5622 | -17.2435
9.602 | -34.5627 | -17.2532
9.6081 | -34.5632 | -17.2631
9.6142 | -34.5637 | -17.2729
9.6203 | -34.5641 | -17.2827
9.6264 | -34.5646 | -17.2925
9.6325 | -34.565 | -17.3024
9.6386 | -34.5655 | -17.3122
9.6447 | -34.566 | -17.322
9.6508 | -34.5665 | -17.3318
9.6569 | -34.5669 | -17.3417
9.663 | -34.5674 | -17.3516
9.6691 | -34.5679 | -17.3615
9.6753 | -34.5684 | -17.3713
9.6814 | -34.5689 | -17.381
9.6875 | -34.5693 | -17.3905
9.6936 | -34.5698 | -17.4003
9.6997 | -34.5703 | -17.4102
9.7058 | -34.5708 | -17.4201
9.7119 | -34.5713 | -17.43
9.718 | -34.5718 | -17.4399
9.7241 | -34.5722 | -17.4496
9.7302 | -34.5727 | -17.4595
9.7363 | -34.5732 | -17.4692
9.7424 | -34.5737 | -17.4789
9.7485 | -34.5741 | -17.4888
9.7546 | -34.5746 | -17.4984
9.7607 | -34.575 | -17.5084
9.7668 | -34.5755 | -17.5182
9.7729 | -34.576 | -17.5281
9.779 | -34.5765 | -17.5377
9.7851 | -34.577 | -17.5476
9.7912 | -34.5775 | -17.5573
9.7973 | -34.578 | -17.5673
9.8034 | -34.5785 | -17.5769
9.8095 | -34.579 | -17.5869
9.8157 | -34.5795 | -17.5967
9.8218 | -34.58 | -17.6067
9.8279 | -34.5805 | -17.6164
9.834 | -34.5809 | -17.6264
9.8401 | -34.5814 | -17.636
9.8462 | -34.5819 | -17.6456
9.8523 | -34.5824 | -17.6554
9.8584 | -34.5829 | -17.6651
9.8645 | -34.5834 | -17.6749
9.8706 | -34.5838 | -17.6848
9.8767 | -34.5843 | -17.6944
9.8828 | -34.5847 | -17.704
9.8889 | -34.5852 | -17.7137
9.895 | -34.5857 | -17.7235
9.9011 | -34.5862 | -17.7333
9.9072 | -34.5866 | -17.743
9.9133 | -34.5871 | -17.7528
9.9194 | -34.5876 | -17.7625
9.9255 | -34.5881 | -17.7722
9.9316 | -34.5885 | -17.7819
9.9377 | -34.589 | -17.7917
9.9438 | -34.5895 | -17.8015
9.9499 | -34.59 | -17.8116
9.956 | -34.5905 | -17.8213
9.9622 | -34.591 | -17.8312
9.9683 | -34.5914 | -17.8409
9.9744 | -34.5919 | -17.8508
9.9805 | -34.5924 | -17.8606
9.9866 | -34.5928 | -17.8703
9.9927 | -34.5933 | -17.88
9.9988 | -34.5938 | -17.8898
10.0049 | -34.5943 | -17.8996
10.011 | -34.5948 | -17.9095
10.0171 | -34.5952 | -17.9191
10.0232 | -34.5957 | -17.9288
10.0293 | -34.5962 | -17.9386
10.0354 | -34.5968 | -17.9484
10.0415 | -34.5973 | -17.9581
10.0476 | -34.5977 | -17.968
10.0537 | -34.5982 | -17.9778
10.0598 | -34.5987 | -17.9875
10.0659 | -34.5991 | -17.9973
10.072 | -34.5996 | -18.0068
10.0781 | -34.6001 | -18.0165
10.0842 | -34.6006 | -18.0263
10.0903 | -34.6011 | -18.0361
10.0964 | -34.6016 | -18.0456
10.1026 | -34.602 | -18.0554
10.1087 | -34.6025 | -18.0651
10.1148 | -34.6029 | -18.075
10.1209 | -34.6034 | -18.0846
10.127 | -34.6039 | -18.0942
10.1331 | -34.6044 | -18.1038
10.1392 | -34.6049 | -18.1134
10.1453 | -34.6054 | -18.1231
10.1514 | -34.6058 | -18.1328
10.1575 | -34.6063 | -18.1427
10.1636 | -34.6068 | -18.1521
10.1697 | -34.6073 | -18.162
10.1758 | -34.6078 | -18.1714
10.1819 | -34.6082 | -18.1811
10.188 | -34.6087 | -18.1908
10.1941 | -34.6092 | -18.2005
10.2002 | -34.6096 | -18.2102
10.2063 | -34.6101 | -18.2198
10.2124 | -34.6106 | -18.2295
10.2185 | -34.6111 | -18.2392
10.2246 | -34.6115 | -18.2488
10.2307 | -34.612 | -18.2586
10.2368 | -34.6125 | -18.2679
10.2429 | -34.613 | -18.2776
10.2491 | -34.6135 | -18.2873
10.2552 | -34.6139 | -18.297
10.2613 | -34.6144 | -18.3067
10.2674 | -34.6149 | -18.3163
10.2735 | -34.6154 | -18.3262
10.2796 | -34.6159 | -18.3358
10.2857 | -34.6164 | -18.3457
10.2918 | -34.6168 | -18.3555
10.2979 | -34.6173 | -18.3651
10.304 | -34.6178 | -18.3744
10.3101 | -34.6183 | -18.3841
10.3162 | -34.6188 | -18.3939
10.3223 | -34.6192 | -18.4033
10.3284 | -34.6197 | -18.4129
10.3345 | -34.6202 | -18.4224
10.3406 | -34.6207 | -18.4319
10.3467 | -34.6211 | -18.4416
10.3528 | -34.6216 | -18.4511
10.3589 | -34.6221 | -18.4607
10.365 | -34.6225 | -18.4704
10.3711 | -34.623 | -18.4801
10.3772 | -34.6235 | -18.4897
10.3833 | -34.624 | -18.4994
10.3895 | -34.6245 | -18.509
10.3956 | -34.625 | -18.5187
10.4017 | -34.6255 | -18.5283
10.4078 | -34.626 | -18.5379
10.4139 | -34.6265 | -18.5477
10.42 | -34.627 | -18.5574
10.4261 | -34.6275 | -18.567
10.4322 | -34.6279 | -18.5766
10.4383 | -34.6284 | -18.5863
10.4444 | -34.6289 | -18.5957
10.4505 | -34.6293 | -18.6054
10.4566 | -34.6298 | -18.6149
10.4627 | -34.6303 | -18.6244
10.4688 | -34.6308 | -18.634
10.4749 | -34.6312 | -18.6436
10.481 | -34.6317 | -18.6533
10.4871 | -34.6322 | -18.6628
10.4932 | -34.6327 | -18.6724
10.4993 | -34.6332 | -18.682
10.5054 | -34.6336 | -18.6918
10.5115 | -34.6341 | -18.7013
10.5176 | -34.6346 | -18.7109
10.5237 | -34.6351 | -18.7206
10.5298 | -34.6356 | -18.73
10.536 | -34.6361 | -18.7396
10.5421 | -34.6366 | -18.7492
10.5482 | -34.637 | -18.7587
10.5543 | -34.6375 | -18.768
10.5604 | -34.638 | -18.7774
10.5665 | -34.6385 | -18.7868
10.5726 | -34.6389 | -18.7964
10.5787 | -34.6394 | -18.8059
10.5848 | -34.6399 | -18.8152
10.5909 | -34.6404 | -18.8248
10.597 | -34.6409 | -18.8345
10.6031 | -34.6414 | -18.8443
10.6092 | -34.6418 | -18.8539
10.6153 | -34.6423 | -18.8636
10.6214 | -34.6428 | -18.8729
10.6275 | -34.6432 | -18.8826
10.6336 | -34.6438 | -18.8923
10.6397 | -34.6442 | -18.902
10.6458 | -34.6447 | -18.9115
10.6519 | -34.6452 | -18.9213
10.658 | -34.6456 | -18.9307
10.6641 | -34.6461 | -18.9404
10.6702 | -34.6466 | -18.9501
10.6764 | -34.6471 | -18.9597
10.6825 | -34.6476 | -18.9693
10.6886 | -34.6481 | -18.9786
10.6947 | -34.6486 | -18.9881
10.7008 | -34.6491 | -18.9976
10.7069 | -34.6496 | -19.0072
10.713 | -34.65 | -19.0168
10.7191 | -34.6505 | -19.0264
10.7252 | -34.651 | -19.0356
10.7313 | -34.6514 | -19.0451
10.7374 | -34.6518 | -19.0546
10.7435 | -34.6523 | -19.0642
10.7496 | -34.6528 | -19.0735
10.7557 | -34.6533 | -19.0831
10.7618 | -34.6538 | -19.092
10.7679 | -34.6542 | -19.1015
10.774 | -34.6547 | -19.1111
10.7801 | -34.6553 | -19.1205
10.7862 | -34.6557 | -19.1297
10.7923 | -34.6562 | -19.1393
10.7984 | -34.6567 | -19.1487
10.8045 | -34.6572 | -19.1578
10.8106 | -34.6576 | -19.1675
10.8168 | -34.6581 | -19.177
10.8229 | -34.6586 | -19.1865
10.829 | -34.659 | -19.1958
10.8351 | -34.6595 | -19.205
10.8412 | -34.66 | -19.2144
10.8473 | -34.6606 | -19.2237
10.8534 | -34.661 | -19.2332
10.8595 | -34.6615 | -19.2426
10.8656 | -34.662 | -19.2521
10.8717 | -34.6625 | -19.2616
10.8778 | -34.663 | -19.2713
10.8839 | -34.6634 | -19.2809
10.89 | -34.6639 | -19.2904
10.8961 | -34.6644 | -19.3
10.9022 | -34.6648 | -19.3094
10.9083 | -34.6653 | -19.3189
10.9144 | -34.6658 | -19.3282
10.9205 | -34.6662 | -19.3377
10.9266 | -34.6667 | -19.347
10.9327 | -34.6672 | -19.3565
10.9388 | -34.6678 | -19.366
10.9449 | -34.6682 | -19.3754
10.951 | -34.6688 | -19.385
10.9571 | -34.6692 | -19.3947
10.9633 | -34.6697 | -19.4041
10.9694 | -34.6702 | -19.4135
10.9755 | -34.6707 | -19.4227
10.9816 | -34.6711 | -19.4322
10.9877 | -34.6716 | -19.4414
10.9938 | -34.6721 | -19.451
10.9999 | -34.6726 | -19.4603
11.006 | -34.6731 | -19.4695
11.0121 | -34.6736 | -19.4789
11.0182 | -34.6741 | -19.4883
11.0243 | -34.6746 | -19.4975
11.0304 | -34.675 | -19.5072
11.0365 | -34.6755 | -19.5166
11.0426 | -34.676 | -19.5261
11.0487 | -34.6765 | -19.5353
11.0548 | -34.677 | -19.5445
11.0609 | -34.6774 | -19.5541
11.067 | -34.6779 | -19.5636
11.0731 | -34.6784 | -19.5731
11.0792 | -34.6789 | -19.5822
11.0853 | -34.6794 | -19.5917
11.0914 | -34.6799 | -19.6012
11.0975 | -34.6804 | -19.6106
11.1037 | -34.6808 | -19.6199
11.1098 | -34.6813 | -19.6295
11.1159 | -34.6817 | -19.6385
11.122 | -34.6822 | -19.6476
11.1281 | -34.6827 | -19.657
11.1342 | -34.6832 | -19.6665
11.1403 | -34.6837 | -19.6759
11.1464 | -34.6841 | -19.6853
11.1525 | -34.6846 | -19.6944
11.1586 | -34.6851 | -19.7036
11.1647 | -34.6856 | -19.7128
11.1708 | -34.6861 | -19.7224
11.1769 | -34.6866 | -19.7318
11.183 | -34.687 | -19.741
11.1891 | -34.6875 | -19.7505
11.1952 | -34.688 | -19.7598
11.2013 | -34.6885 | -19.7691
11.2074 | -34.689 | -19.7784
11.2135 | -34.6895 | -19.7879
11.2196 | -34.69 | -19.7973
11.2257 | -34.6905 | -19.8068
11.2318 | -34.691 | -19.8162
11.2379 | -34.6915 | -19.8256
11.244 | -34.692 | -19.835
11.2502 | -34.6925 | -19.8442
11.2563 | -34.693 | -19.8537
11.2624 | -34.6934 | -19.8629
11.2685 | -34.6939 | -19.8724
11.2746 | -34.6944 | -19.8814
11.2807 | -34.6949 | -19.8906
11.2868 | -34.6954 | -19.9
11.2929 | -34.6958 | -19.9091
11.299 | -34.6964 | -19.9186
11.3051 | -34.6968 | -19.9282
11.3112 | -34.6973 | -19.9374
11.3173 | -34.6978 | -19.9464
11.3234 | -34.6983 | -19.9556
11.3295 | -34.6988 | -19.965
11.3356 | -34.6993 | -19.9743
11.3417 | -34.6998 | -19.9836
11.3478 | -34.7003 | -19.9933
11.3539 | -34.7007 | -20.0027
11.36 | -34.7013 | -20.0119
11.3661 | -34.7018 | -20.0211
11.3722 | -34.7023 | -20.03
11.3783 | -34.7027 | -20.0395
11.3844 | -34.7032 | -20.0489
11.3906 | -34.7037 | -20.0584
11.3967 | -34.7042 | -20.0677
11.4028 | -34.7047 | -20.0772
11.4089 | -34.7052 | -20.0863
11.415 | -34.7057 | -20.0954
11.4211 | -34.7062 | -20.1044
11.4272 | -34.7067 | -20.1137
11.4333 | -34.7073 | -20.1228
11.4394 | -34.7077 | -20.1321
11.4455 | -34.7082 | -20.1412
11.4516 | -34.7087 | -20.1503
11.4577 | -34.7091 | -20.1596
11.4638 | -34.7097 | -20.1689
11.4699 | -34.7101 | -20.1782
11.476 | -34.7106 | -20.1878
11.4821 | -34.711 | -20.1968
11.4882 | -34.7116 | -20.2061
11.4943 | -34.712 | -20.2155
11.5004 | -34.7125 | -20.2246
11.5065 | -34.713 | -20.2338
11.5126 | -34.7135 | -20.2431
11.5187 | -34.714 | -20.2525
11.5248 | -34.7145 | -20.262
11.5309 | -34.715 | -20.2712
11.5371 | -34.7154 | -20.2802
11.5432 | -34.7159 | -20.2897
11.5493 | -34.7164 | -20.2991
11.5554 | -34.7169 | -20.3084
11.5615 | -34.7175 | -20.3178
11.5676 | -34.718 | -20.3272
11.5737 | -34.7185 | -20.3365
11.5798 | -34.719 | -20.3456
11.5859 | -34.7195 | -20.355
11.592 | -34.7199 | -20.3641
11.5981 | -34.7204 | -20.3735
11.6042 | -34.721 | -20.3826
11.6103 | -34.7215 | -20.3917
11.6164 | -34.722 | -20.4009
11.6225 | -34.7224 | -20.4104
11.6286 | -34.7229 | -20.4196
11.6347 | -34.7234 | -20.4289
11.6408 | -34.7239 | -20.438
11.6469 | -34.7244 | -20.4472
11.653 | -34.7249 | -20.4563
11.6591 | -34.7254 | -20.4658
11.6652 | -34.7259 | -20.475
11.6713 | -34.7264 | -20.4841
11.6775 | -34.7269 | -20.493
11.6836 | -34.7274 | -20.5022
11.6897 | -34.7279 | -20.5114
11.6958 | -34.7284 | -20.5207
11.7019 | -34.7289 | -20.53
11.708 | -34.7294 | -20.5392
11.7141 | -34.7299 | -20.548
11.7202 | -34.7304 | -20.5574
11.7263 | -34.7309 | -20.5667
11.7324 | -34.7314 | -20.5762
11.7385 | -34.7318 | -20.5854
11.7446 | -34.7323 | -20.5948
11.7507 | -34.7328 | -20.604
11.7568 | -34.7333 | -20.6133
11.7629 | -34.7338 | -20.6225
11.769 | -34.7343 | -20.6316
11.7751 | -34.7348 | -20.6408
11.7812 | -34.7353 | -20.6498
11.7873 | -34.7358 | -20.659
11.7934 | -34.7363 | -20.6684
11.7995 | -34.7368 | -20.6776
11.8056 | -34.7373 | -20.6867
11.8117 | -34.7378 | -20.696
11.8178 | -34.7383 | -20.7052
11.824 | -34.7388 | -20.7143
11.8301 | -34.7393 | -20.7234
11.8362 | -34.7398 | -20.7325
11.8423 | -34.7403 | -20.7417
11.8484 | -34.7408 | -20.7509
11.8545 | -34.7414 | -20.7599
11.8606 | -34.7419 | -20.7689
11.8667 | -34.7424 | -20.7783
11.8728 | -34.7429 | -20.7871
11.8789 | -34.7433 | -20.7962
11.885 | -34.7439 | -20.8056
11.8911 | -34.7444 | -20.8148
11.8972 | -34.7449 | -20.8239
11.9033 | -34.7454 | -20.8331
11.9094 | -34.7459 | -20.8423
11.9155 | -34.7464 | -20.8515
11.9216 | -34.7469 | -20.8605
11.9277 | -34.7474 | -20.8695
11.9338 | -34.7479 | -20.8787
11.9399 | -34.7484 | -20.888
11.946 | -34.7489 | -20.897
11.9521 | -34.7493 | -20.9061
11.9582 | -34.7499 | -20.9152
11.9644 | -34.7504 | -20.9241
11.9705 | -34.7509 | -20.9332
11.9766 | -34.7514 | -20.9422
11.9827 | -34.7519 | -20.9514
11.9888 | -34.7525 | -20.9605
11.9949 | -34.753 | -20.9693
12.001 | -34.7535 | -20.9785
12.0071 | -34.754 | -20.9876
12.0132 | -34.7545 | -20.9966
12.0193 | -34.7549 | -21.006
12.0254 | -34.7554 | -21.0152
12.0315 | -34.756 | -21.0241
12.0376 | -34.7565 | -21.0331
12.0437 | -34.7569 | -21.0423
12.0498 | -34.7575 | -21.0515
12.0559 | -34.758 | -21.0606
12.062 | -34.7585 | -21.0694
12.0681 | -34.759 | -21.0786
12.0742 | -34.7595 | -21.0876
12.0803 | -34.76 | -21.0969
12.0864 | -34.7605 | -21.106
12.0925 | -34.761 | -21.1151
12.0986 | -34.7616 | -21.124
12.1047 | -34.7621 | -21.1331
12.1109 | -34.7626 | -21.142
12.117 | -34.7631 | -21.1511
12.1231 | -34.7636 | -21.1602
12.1292 | -34.7641 | -21.1691
12.1353 | -34.7646 | -21.1783
12.1414 | -34.7651 | -21.1874
12.1475 | -34.7656 | -21.1966
12.1536 | -34.7661 | -21.2056
12.1597 | -34.7666 | -21.2146
12.1658 | -34.7671 | -21.2236
12.1719 | -34.7677 | -21.2325
12.178 | -34.7682 | -21.2416
12.1841 | -34.7687 | -21.2507
12.1902 | -34.7692 | -21.2598
12.1963 | -34.7697 | -21.2689
12.2024 | -34.7703 | -21.2779
12.2085 | -34.7708 | -21.2869
12.2146 | -34.7713 | -21.296
12.2207 | -34.7718 | -21.305
12.2268 | -34.7723 | -21.314
12.2329 | -34.7728 | -21.3229
12.239 | -34.7734 | -21.3318
12.2451 | -34.7739 | -21.3408
12.2513 | -34.7744 | -21.3499
12.2574 | -34.7749 | -21.3589
12.2635 | -34.7754 | -21.368
12.2696 | -34.7759 | -21.3769
12.2757 | -34.7765 | -21.3858
12.2818 | -34.777 | -21.395
12.2879 | -34.7775 | -21.4039
12.294 | -34.778 | -21.4129
12.3001 | -34.7785 | -21.4219
12.3062 | -34.7791 | -21.431
12.3123 | -34.7796 | -21.4398
12.3184 | -34.7802 | -21.4486
12.3245 | -34.7807 | -21.4575
12.3306 | -34.7812 | -21.4665
12.3367 | -34.7817 | -21.4755
12.3428 | -34.7822 | -21.4844
12.3489 | -34.7828 | -21.4932
12.355 | -34.7833 | -21.5024
12.3611 | -34.7838 | -21.5116
12.3672 | -34.7843 | -21.5205
12.3733 | -34.7848 | -21.5296
12.3794 | -34.7853 | -21.5383
12.3855 | -34.7858 | -21.5473
12.3916 | -34.7863 | -21.5563
12.3978 | -34.7869 | -21.5654
12.4039 | -34.7874 | -21.5745
12.41 | -34.7879 | -21.5835
12.4161 | -34.7885 | -21.592
12.4222 | -34.789 | -21.601
12.4283 | -34.7895 | -21.6099
12.4344 | -34.79 | -21.619
12.4405 | -34.7906 | -21.6281
12.4466 | -34.7911 | -21.6371
12.4527 | -34.7916 | -21.6461
12.4588 | -34.7921 | -21.655
12.4649 | -34.7926 | -21.664
12.471 | -34.7932 | -21.6731
12.4771 | -34.7937 | -21.6821
12.4832 | -34.7942 | -21.6911
12.4893 | -34.7947 | -21.6998
12.4954 | -34.7952 | -21.7088
12.5015 | -34.7957 | -21.7177
12.5076 | -34.7962 | -21.7267
12.5137 | -34.7968 | -21.7357
12.5198 | -34.7973 | -21.7445
12.5259 | -34.7978 | -21.7534
12.532 | -34.7984 | -21.7624
12.5382 | -34.7989 | -21.7714
12.5443 | -34.7994 | -21.7803
12.5504 | -34.7999 | -21.7892
12.5565 | -34.8004 | -21.798
12.5626 | -34.801 | -21.8071
12.5687 | -34.8015 | -21.8161
12.5748 | -34.802 | -21.8251
12.5809 | -34.8025 | -21.8341
12.587 | -34.803 | -21.8432
12.5931 | -34.8036 | -21.8521
12.5992 | -34.8042 | -21.861
12.6053 | -34.8047 | -21.87
12.6114 | -34.8052 | -21.8789
12.6175 | -34.8057 | -21.888
12.6236 | -34.8062 | -21.8971
12.6297 | -34.8068 | -21.9059
12.6358 | -34.8073 | -21.9145
12.6419 | -34.8079 | -21.9235
12.648 | -34.8084 | -21.9324
12.6541 | -34.8089 | -21.9413
12.6602 | -34.8095 | -21.95
12.6663 | -34.81 | -21.9589
12.6724 | -34.8105 | -21.9678
12.6785 | -34.811 | -21.9766
12.6847 | -34.8116 | -21.9855
12.6908 | -34.8121 | -21.9944
12.6969 | -34.8127 | -22.0033
12.703 | -34.8132 | -22.0122
12.7091 | -34.8137 | -22.0211
12.7152 | -34.8142 | -22.0299
12.7213 | -34.8147 | -22.0389
12.7274 | -34.8153 | -22.0479
12.7335 | -34.8158 | -22.0567
12.7396 | -34.8164 | -22.0653
12.7457 | -34.8169 | -22.0743
12.7518 | -34.8174 | -22.0834
12.7579 | -34.818 | -22.0923
12.764 | -34.8185 | -22.1013
12.7701 | -34.819 | -22.1104
12.7762 | -34.8196 | -22.1191
12.7823 | -34.8201 | -22.1281
12.7884 | -34.8207 | -22.1371
12.7945 | -34.8212 | -22.1462
12.8006 | -34.8218 | -22.155
12.8067 | -34.8223 | -22.1638
12.8128 | -34.8228 | -22.1725
12.8189 | -34.8233 | -22.1814
12.8251 | -34.8239 | -22.1904
12.8312 | -34.8244 | -22.1994
12.8373 | -34.8249 | -22.2082
12.8434 | -34.8255 | -22.2173
12.8495 | -34.8261 | -22.226
12.8556 | -34.8266 | -22.2349
12.8617 | -34.8271 | -22.2437
12.8678 | -34.8276 | -22.2529
12.8739 | -34.8281 | -22.2617
12.88 | -34.8287 | -22.2706
12.8861 | -34.8292 | -22.2793
12.8922 | -34.8297 | -22.2883
12.8983 | -34.8303 | -22.2972
12.9044 | -34.8308 | -22.306
12.9105 | -34.8314 | -22.3147
12.9166 | -34.8319 | -22.3236
12.9227 | -34.8324 | -22.3325
12.9288 | -34.833 | -22.3413
12.9349 | -34.8335 | -22.3502
12.941 | -34.8341 | -22.3591
12.9471 | -34.8346 | -22.368
12.9532 | -34.8352 | -22.3769
12.9593 | -34.8357 | -22.386
12.9654 | -34.8362 | -22.395
12.9716 | -34.8368 | -22.404
12.9777 | -34.8373 | -22.4128
12.9838 | -34.8379 | -22.4218
12.9899 | -34.8384 | -22.4308
12.996 | -34.8389 | -22.4397
13.0021 | -34.8395 | -22.4486
13.0082 | -34.84 | -22.4574
13.0143 | -34.8406 | -22.466
13.0204 | -34.8412 | -22.4746
13.0265 | -34.8417 | -22.4832
13.0326 | -34.8423 | -22.492
13.0387 | -34.8428 | -22.5008
13.0448 | -34.8433 | -22.5096
13.0509 | -34.8439 | -22.5185
13.057 | -34.8444 | -22.5274
13.0631 | -34.8449 | -22.5363
13.0692 | -34.8455 | -22.5453
13.0753 | -34.846 | -22.5543
13.0814 | -34.8466 | -22.5632
13.0875 | -34.8472 | -22.5718
13.0936 | -34.8477 | -22.5806
13.0997 | -34.8482 | -22.5893
13.1058 | -34.8488 | -22.5979
13.112 | -34.8493 | -22.6067
13.1181 | -34.8499 | -22.6154
13.1242 | -34.8504 | -22.6242
13.1303 | -34.851 | -22.6331
13.1364 | -34.8515 | -22.642
13.1425 | -34.852 | -22.6506
13.1486 | -34.8526 | -22.6595
13.1547 | -34.8531 | -22.6682
13.1608 | -34.8536 | -22.6769
13.1669 | -34.8541 | -22.6855
13.173 | -34.8547 | -22.6943
13.1791 | -34.8552 | -22.7031
13.1852 | -34.8558 | -22.7117
13.1913 | -34.8563 | -22.7205
13.1974 | -34.8568 | -22.7293
13.2035 | -34.8573 | -22.738
13.2096 | -34.8579 | -22.7467
13.2157 | -34.8584 | -22.7554
13.2218 | -34.859 | -22.7641
13.2279 | -34.8595 | -22.7728
13.234 | -34.8601 | -22.7815
13.2401 | -34.8606 | -22.7903
13.2462 | -34.8611 | -22.799
13.2524 | -34.8617 | -22.8079
13.2585 | -34.8623 | -22.8165
13.2646 | -34.8629 | -22.8256
13.2707 | -34.8634 | -22.8343
13.2768 | -34.864 | -22.8433
13.2829 | -34.8645 | -22.8522
13.289 | -34.8651 | -22.8609
13.2951 | -34.8656 | -22.8692
13.3012 | -34.8661 | -22.878
13.3073 | -34.8667 | -22.8869
13.3134 | -34.8672 | -22.8958
13.3195 | -34.8678 | -22.9045
13.3256 | -34.8683 | -22.9132
13.3317 | -34.8688 | -22.922
13.3378 | -34.8693 | -22.9309
13.3439 | -34.8699 | -22.9398
13.35 | -34.8704 | -22.9487
13.3561 | -34.8709 | -22.9572
13.3622 | -34.8714 | -22.9662
13.3683 | -34.872 | -22.9748
13.3744 | -34.8725 | -22.9834
13.3805 | -34.8731 | -22.9921
13.3866 | -34.8736 | -23.0009
13.3927 | -34.8742 | -23.0098
13.3989 | -34.8747 | -23.0184
13.405 | -34.8752 | -23.0273
13.4111 | -34.8757 | -23.036
13.4172 | -34.8763 | -23.0448
13.4233 | -34.8768 | -23.0537
13.4294 | -34.8773 | -23.0624
13.4355 | -34.8779 | -23.0711
13.4416 | -34.8784 | -23.0795
13.4477 | -34.879 | -23.0883
13.4538 | -34.8795 | -23.0969
13.4599 | -34.88 | -23.1058
13.466 | -34.8806 | -23.1146
13.4721 | -34.8812 | -23.1233
13.4782 | -34.8817 | -23.1321
13.4843 | -34.8823 | -23.1407
13.4904 | -34.8828 | -23.1497
13.4965 | -34.8834 | -23.1584
13.5026 | -34.8839 | -23.1673
13.5087 | -34.8844 | -23.1759
13.5148 | -34.885 | -23.1848
13.5209 | -34.8855 | -23.1937
13.527 | -34.886 | -23.2022
13.5331 | -34.8866 | -23.2109
13.5393 | -34.8871 | -23.2192
13.5454 | -34.8877 | -23.2277
13.5515 | -34.8882 | -23.2367
13.5576 | -34.8887 | -23.2456
13.5637 | -34.8893 | -23.2544
13.5698 | -34.8898 | -23.2632
13.5759 | -34.8904 | -23.2717
13.582 | -34.8909 | -23.2803
13.5881 | -34.8915 | -23.2889
13.5942 | -34.892 | -23.2979
13.6003 | -34.8925 | -23.3067
13.6064 | -34.8931 | -23.3156
13.6125 | -34.8937 | -23.3242
13.6186 | -34.8942 | -23.333
13.6247 | -34.8948 | -23.3418
13.6308 | -34.8953 | -23.3507
13.6369 | -34.8958 | -23.3595
13.643 | -34.8964 | -23.3685
13.6491 | -34.8969 | -23.3771
13.6552 | -34.8974 | -23.386
13.6613 | -34.898 | -23.3946
13.6674 | -34.8985 | -23.4034
13.6735 | -34.899 | -23.4123
13.6796 | -34.8996 | -23.4212
13.6858 | -34.9001 | -23.4297
13.6919 | -34.9006 | -23.4384
13.698 | -34.9012 | -23.4471
13.7041 | -34.9017 | -23.4561
13.7102 | -34.9023 | -23.4649
13.7163 | -34.9029 | -23.4734
13.7224 | -34.9034 | -23.4822
13.7285 | -34.904 | -23.4908
13.7346 | -34.9045 | -23.4996
13.7407 | -34.905 | -23.5084
13.7468 | -34.9056 | -23.517
13.7529 | -34.9061 | -23.5255
13.759 | -34.9067 | -23.5343
13.7651 | -34.9072 | -23.5432
13.7712 | -34.9077 | -23.5519
13.7773 | -34.9083 | -23.5606
13.7834 | -34.9088 | -23.5694
13.7895 | -34.9094 | -23.578
13.7956 | -34.9099 | -23.5869
13.8017 | -34.9104 | -23.5955
13.8078 | -34.9109 | -23.6042
13.8139 | -34.9115 | -23.6129
13.82 | -34.912 | -23.6215
13.8262 | -34.9125 | -23.6302
13.8323 | -34.9131 | -23.6389
13.8384 | -34.9136 | -23.6477
13.8445 | -34.9142 | -23.6565
13.8506 | -34.9147 | -23.6651
13.8567 | -34.9153 | -23.6739
13.8628 | -34.9158 | -23.6827
13.8689 | -34.9164 | -23.6915
13.875 | -34.9169 | -23.7002
13.8811 | -34.9175 | -23.709
13.8872 | -34.918 | -23.7177
13.8933 | -34.9185 | -23.7261
13.8994 | -34.919 | -23.7349
13.9055 | -34.9196 | -23.7437
13.9116 | -34.9201 | -23.7524
13.9177 | -34.9206 | -23.7612
13.9238 | -34.9211 | -23.7697
13.9299 | -34.9217 | -23.7783
13.936 | -34.9222 | -23.787
13.9421 | -34.9227 | -23.7956
13.9482 | -34.9232 | -23.8042
13.9543 | -34.9238 | -23.813
13.9604 | -34.9243 | -23.8216
13.9665 | -34.9249 | -23.8302
13.9727 | -34.9254 | -23.8391
13.9788 | -34.926 | -23.8477
13.9849 | -34.9265 | -23.8565
13.991 | -34.927 | -23.8652
13.9971 | -34.9275 | -23.8735
14.0032 | -34.928 | -23.8823
14.0093 | -34.9286 | -23.891
14.0154 | -34.9291 | -23.8998
14.0215 | -34.9296 | -23.9084
14.0276 | -34.9302 | -23.9171
14.0337 | -34.9307 | -23.9258
14.0398 | -34.9313 | -23.9345
14.0459 | -34.9318 | -23.943
14.052 | -34.9324 | -23.9519
14.0581 | -34.9329 | -23.9607
14.0642 | -34.9334 | -23.9693
14.0703 | -34.9339 | -23.978
14.0764 | -34.9345 | -23.9865
14.0825 | -34.935 | -23.9951
14.0886 | -34.9356 | -24.0038
14.0947 | -34.9361 | -24.0126
14.1008 | -34.9367 | -24.0212
14.1069 | -34.9372 | -24.0295
14.1131 | -34.9377 | -24.0382
14.1192 | -34.9383 | -24.0468
14.1253 | -34.9388 | -24.0553
14.1314 | -34.9394 | -24.0639
14.1375 | -34.9399 | -24.0726
14.1436 | -34.9405 | -24.0812
14.1497 | -34.941 | -24.0898
14.1558 | -34.9415 | -24.0984
14.1619 | -34.942 | -24.107
14.168 | -34.9426 | -24.1157
14.1741 | -34.9432 | -24.124
14.1802 | -34.9437 | -24.1328
14.1863 | -34.9442 | -24.1414
14.1924 | -34.9448 | -24.1502
14.1985 | -34.9453 | -24.1586
14.2046 | -34.9459 | -24.1674
14.2107 | -34.9464 | -24.1757
14.2168 | -34.947 | -24.1844
14.2229 | -34.9475 | -24.1929
14.229 | -34.948 | -24.2015
14.2351 | -34.9486 | -24.2101
14.2412 | -34.9491 | -24.2187
14.2473 | -34.9497 | -24.2272
14.2534 | -34.9502 | -24.2356
14.2596 | -34.9507 | -24.2442
14.2657 | -34.9512 | -24.2528
14.2718 | -34.9518 | -24.2614
14.2779 | -34.9523 | -24.2701
14.284 | -34.9528 | -24.2787
14.2901 | -34.9534 | -24.2873
14.2962 | -34.9539 | -24.296
14.3023 | -34.9545 | -24.3046
14.3084 | -34.955 | -24.3132
14.3145 | -34.9556 | -24.3217
14.3206 | -34.9561 | -24.3304
14.3267 | -34.9567 | -24.339
14.3328 | -34.9572 | -24.3476
14.3389 | -34.9577 | -24.3562
14.345 | -34.9582 | -24.3648
14.3511 | -34.9588 | -24.3734
14.3572 | -34.9593 | -24.3821
14.3633 | -34.9599 | -24.3908
14.3694 | -34.9604 | -24.3994
14.3755 | -34.961 | -24.4078
14.3816 | -34.9615 | -24.4167
14.3877 | -34.9621 | -24.425
14.3938 | -34.9626 | -24.4337
14.4 | -34.9631 | -24.4422
14.4061 | -34.9637 | -24.4509
14.4122 | -34.9642 | -24.4596
14.4183 | -34.9647 | -24.4681
14.4244 | -34.9653 | -24.4766
14.4305 | -34.9658 | -24.4851
14.4366 | -34.9664 | -24.4938
14.4427 | -34.9669 | -24.5025
14.4488 | -34.9674 | -24.5111
14.4549 | -34.968 | -24.5196
14.461 | -34.9686 | -24.5281
14.4671 | -34.9691 | -24.5368
14.4732 | -34.9696 | -24.5455
14.4793 | -34.9702 | -24.5543
14.4854 | -34.9707 | -24.5629
14.4915 | -34.9713 | -24.5712
14.4976 | -34.9719 | -24.5796
14.5037 | -34.9724 | -24.5881
14.5098 | -34.9729 | -24.5968
14.5159 | -34.9734 | -24.6054
14.522 | -34.974 | -24.6139
14.5281 | -34.9745 | -24.6222
14.5342 | -34.975 | -24.6309
14.5403 | -34.9756 | -24.6394
14.5465 | -34.9761 | -24.6479
14.5526 | -34.9767 | -24.6566
14.5587 | -34.9772 | -24.6652
14.5648 | -34.9777 | -24.6734
14.5709 | -34.9783 | -24.6819
14.577 | -34.9788 | -24.6905
14.5831 | -34.9793 | -24.6991
14.5892 | -34.9799 | -24.7076
14.5953 | -34.9804 | -24.7159
14.6014 | -34.981 | -24.7245
14.6075 | -34.9815 | -24.733
14.6136 | -34.982 | -24.7417
14.6197 | -34.9825 | -24.7501
14.6258 | -34.9831 | -24.7587
14.6319 | -34.9836 | -24.767
14.638 | -34.9841 | -24.7756
14.6441 | -34.9846 | -24.7843
14.6502 | -34.9852 | -24.793
14.6563 | -34.9857 | -24.8015
14.6624 | -34.9863 | -24.8102
14.6685 | -34.9868 | -24.8185
14.6746 | -34.9873 | -24.8271
14.6807 | -34.9879 | -24.8355
14.6869 | -34.9884 | -24.8441
14.693 | -34.989 | -24.8525
14.6991 | -34.9895 | -24.861
14.7052 | -34.9901 | -24.8693
14.7113 | -34.9906 | -24.8779
14.7174 | -34.9911 | -24.8865
14.7235 | -34.9917 | -24.8949
14.7296 | -34.9922 | -24.9033
14.7357 | -34.9928 | -24.9118
14.7418 | -34.9933 | -24.9203
14.7479 | -34.9939 | -24.9288
14.754 | -34.9944 | -24.9373
14.7601 | -34.995 | -24.9457
14.7662 | -34.9955 | -24.9542
14.7723 | -34.996 | -24.9624
14.7784 | -34.9965 | -24.9709
14.7845 | -34.9971 | -24.9795
14.7906 | -34.9976 | -24.9881
14.7967 | -34.9981 | -24.9966
14.8028 | -34.9987 | -25.0049
14.8089 | -34.9992 | -25.0133
14.815 | -34.9998 | -25.0219
14.8211 | -35.0003 | -25.0304
14.8272 | -35.0009 | -25.0389
14.8334 | -35.0014 | -25.0474
14.8395 | -35.002 | -25.0558
14.8456 | -35.0026 | -25.0641
14.8517 | -35.0031 | -25.0726
14.8578 | -35.0036 | -25.0812
14.8639 | -35.0042 | -25.0898
14.87 | -35.0047 | -25.0983
14.8761 | -35.0053 | -25.1066
14.8822 | -35.0059 | -25.1151
14.8883 | -35.0064 | -25.1236
14.8944 | -35.0069 | -25.1318
14.9005 | -35.0075 | -25.1403
14.9066 | -35.008 | -25.1487
14.9127 | -35.0086 | -25.157
14.9188 | -35.0091 | -25.1655
14.9249 | -35.0096 | -25.1741
14.931 | -35.0102 | -25.1825
14.9371 | -35.0107 | -25.191
14.9432 | -35.0113 | -25.1994
14.9493 | -35.0118 | -25.2078
14.9554 | -35.0124 | -25.2162
14.9615 | -35.0129 | -25.2247
14.9676 | -35.0134 | -25.2334
14.9738 | -35.014 | -25.2418
14.9799 | -35.0145 | -25.2502
14.986 | -35.0151 | -25.2585
14.9921 | -35.0156 | -25.267
14.9982 | -35.0161 | -25.2756
15.0043 | -35.0166 | -25.2839
15.0104 | -35.0172 | -25.2924
15.0165 | -35.0177 | -25.301
15.0226 | -35.0183 | -25.3094
15.0287 | -35.0188 | -25.3179
15.0348 | -35.0194 | -25.3264
15.0409 | -35.0199 | -25.3348
15.047 | -35.0204 | -25.3433
15.0531 | -35.021 | -25.3517
15.0592 | -35.0215 | -25.3602
15.0653 | -35.0221 | -25.3686
15.0714 | -35.0227 | -25.3769
15.0775 | -35.0232 | -25.3853
15.0836 | -35.0238 | -25.3937
15.0897 | -35.0244 | -25.402
15.0958 | -35.0249 | -25.4104
15.1019 | -35.0255 | -25.4188
15.108 | -35.026 | -25.4271
15.1141 | -35.0265 | -25.4354
15.1203 | -35.0271 | -25.4439
15.1264 | -35.0277 | -25.4522
15.1325 | -35.0282 | -25.4606
15.1386 | -35.0287 | -25.4688
15.1447 | -35.0293 | -25.4772
15.1508 | -35.0299 | -25.4857
15.1569 | -35.0304 | -25.4941
15.163 | -35.031 | -25.5022
15.1691 | -35.0315 | -25.5105
15.1752 | -35.0321 | -25.5189
15.1813 | -35.0326 | -25.5271
15.1874 | -35.0332 | -25.5355
15.1935 | -35.0337 | -25.5437
15.1996 | -35.0343 | -25.5522
15.2057 | -35.0348 | -25.5606
15.2118 | -35.0354 | -25.569
15.2179 | -35.0359 | -25.5775
15.224 | -35.0365 | -25.5859
15.2301 | -35.037 | -25.5942
15.2362 | -35.0376 | -25.6026
15.2423 | -35.0382 | -25.6112
15.2484 | -35.0387 | -25.6195
15.2545 | -35.0393 | -25.6279
15.2607 | -35.0399 | -25.6364
15.2668 | -35.0404 | -25.6445
15.2729 | -35.041 | -25.6529
15.279 | -35.0416 | -25.6616
15.2851 | -35.0421 | -25.6699
15.2912 | -35.0427 | -25.6784
15.2973 | -35.0432 | -25.687
15.3034 | -35.0438 | -25.6952
15.3095 | -35.0443 | -25.7037
15.3156 | -35.0449 | -25.712
15.3217 | -35.0454 | -25.7205
15.3278 | -35.046 | -25.729
15.3339 | -35.0466 | -25.7371
15.34 | -35.0471 | -25.7454
15.3461 | -35.0477 | -25.7538
15.3522 | -35.0482 | -25.7621
15.3583 | -35.0488 | -25.7708
15.3644 | -35.0493 | -25.779
15.3705 | -35.0499 | -25.7871
15.3766 | -35.0504 | -25.7954
15.3827 | -35.051 | -25.8037
15.3888 | -35.0516 | -25.8121
15.3949 | -35.0521 | -25.8205
15.401 | -35.0527 | -25.829
15.4072 | -35.0532 | -25.8371
15.4133 | -35.0538 | -25.8454
15.4194 | -35.0543 | -25.8538
15.4255 | -35.0549 | -25.8622
15.4316 | -35.0554 | -25.8704
15.4377 | -35.056 | -25.8789
15.4438 | -35.0566 | -25.8871
15.4499 | -35.0571 | -25.8951
15.456 | -35.0577 | -25.9033
15.4621 | -35.0583 | -25.9114
15.4682 | -35.0588 | -25.9198
15.4743 | -35.0594 | -25.928
15.4804 | -35.0599 | -25.9363
15.4865 | -35.0605 | -25.9446
15.4926 | -35.0611 | -25.953
15.4987 | -35.0617 | -25.9613
15.5048 | -35.0622 | -25.9695
15.5109 | -35.0628 | -25.9778
15.517 | -35.0633 | -25.9863
15.5231 | -35.0639 | -25.9946
15.5292 | -35.0644 | -26.0029
15.5353 | -35.065 | -26.0113
15.5414 | -35.0655 | -26.0195
15.5476 | -35.0661 | -26.0276
15.5537 | -35.0667 | -26.0357
15.5598 | -35.0672 | -26.0441
15.5659 | -35.0677 | -26.0523
15.572 | -35.0683 | -26.0604
15.5781 | -35.0689 | -26.0689
15.5842 | -35.0695 | -26.0771
15.5903 | -35.07 | -26.0854
15.5964 | -35.0706 | -26.0937
15.6025 | -35.0712 | -26.102
15.6086 | -35.0718 | -26.1103
15.6147 | -35.0723 | -26.1187
15.6208 | -35.0729 | -26.1269
15.6269 | -35.0735 | -26.1351
15.633 | -35.074 | -26.1434
15.6391 | -35.0746 | -26.1517
15.6452 | -35.0751 | -26.16
15.6513 | -35.0757 | -26.1681
15.6574 | -35.0763 | -26.1766
15.6635 | -35.0768 | -26.1847
15.6696 | -35.0774 | -26.1932
15.6757 | -35.078 | -26.2013
15.6818 | -35.0786 | -26.2094
15.688 | -35.0791 | -26.2177
15.6941 | -35.0797 | -26.2258
15.7002 | -35.0803 | -26.2342
15.7063 | -35.0808 | -26.2425
15.7124 | -35.0813 | -26.2509
15.7185 | -35.0819 | -26.2591
15.7246 | -35.0825 | -26.2672
15.7307 | -35.0831 | -26.2754
15.7368 | -35.0837 | -26.2835
15.7429 | -35.0842 | -26.2917
15.749 | -35.0848 | -26.3
15.7551 | -35.0853 | -26.3083
15.7612 | -35.0859 | -26.3164
15.7673 | -35.0865 | -26.3247
15.7734 | -35.0871 | -26.3329
15.7795 | -35.0876 | -26.3411
15.7856 | -35.0882 | -26.3492
15.7917 | -35.0888 | -26.3572
15.7978 | -35.0893 | -26.3656
15.8039 | -35.0899 | -26.374
15.81 | -35.0905 | -26.3823
15.8161 | -35.0911 | -26.3904
15.8222 | -35.0916 | -26.3984
15.8283 | -35.0922 | -26.4062
15.8345 | -35.0928 | -26.4146
15.8406 | -35.0934 | -26.4228
15.8467 | -35.0939 | -26.4312
15.8528 | -35.0945 | -26.4394
15.8589 | -35.0951 | -26.4478
15.865 | -35.0957 | -26.4557
15.8711 | -35.0962 | -26.4641
15.8772 | -35.0968 | -26.4719
15.8833 | -35.0973 | -26.4803
15.8894 | -35.0979 | -26.4886
15.8955 | -35.0985 | -26.4967
15.9016 | -35.0991 | -26.5046
15.9077 | -35.0996 | -26.5127
15.9138 | -35.1002 | -26.5211
15.9199 | -35.1007 | -26.5292
15.926 | -35.1013 | -26.5373
15.9321 | -35.1018 | -26.5453
15.9382 | -35.1024 | -26.5534
15.9443 | -35.103 | -26.5616
15.9504 | -35.1035 | -26.5697
15.9565 | -35.1041 | -26.5778
15.9626 | -35.1046 | -26.586
15.9687 | -35.1052 | -26.5938
15.9749 | -35.1057 | -26.6018
15.981 | -35.1063 | -26.6101
15.9871 | -35.1069 | -26.6185
15.9932 | -35.1074 | -26.6265
15.9993 | -35.108 | -26.6347
16.0054 | -35.1086 | -26.6425
16.0115 | -35.1092 | -26.6507
16.0176 | -35.1097 | -26.6588
16.0237 | -35.1103 | -26.6672
16.0298 | -35.1108 | -26.6754
16.0359 | -35.1114 | -26.6837
16.042 | -35.112 | -26.6916
16.0481 | -35.1126 | -26.6997
16.0542 | -35.1132 | -26.7079
16.0603 | -35.1137 | -26.7159
16.0664 | -35.1142 | -26.7239
16.0725 | -35.1148 | -26.7317
16.0786 | -35.1154 | -26.7399
16.0847 | -35.116 | -26.7483
16.0908 | -35.1165 | -26.7564
16.0969 | -35.1171 | -26.7647
16.103 | -35.1176 | -26.7728
16.1091 | -35.1183 | -26.7807
16.1152 | -35.1188 | -26.7889
16.1214 | -35.1194 | -26.797
16.1275 | -35.12 | -26.805
16.1336 | -35.1205 | -26.8131
16.1397 | -35.1211 | -26.8211
16.1458 | -35.1217 | -26.829
16.1519 | -35.1222 | -26.8372
16.158 | -35.1228 | -26.8453
16.1641 | -35.1234 | -26.8537
16.1702 | -35.124 | -26.8618
16.1763 | -35.1245 | -26.8698
16.1824 | -35.1251 | -26.8776
16.1885 | -35.1257 | -26.8854
16.1946 | -35.1262 | -26.8935
16.2007 | -35.1268 | -26.9015
16.2068 | -35.1274 | -26.9094
16.2129 | -35.128 | -26.9176
16.219 | -35.1286 | -26.9254
16.2251 | -35.1291 | -26.9335
16.2312 | -35.1297 | -26.9419
16.2373 | -35.1303 | -26.9498
16.2434 | -35.1308 | -26.9579
16.2495 | -35.1314 | -26.9656
16.2556 | -35.132 | -26.9738
16.2618 | -35.1326 | -26.9817
16.2679 | -35.1331 | -26.99
16.274 | -35.1337 | -26.9981
16.2801 | -35.1343 | -27.0061
16.2862 | -35.1349 | -27.014
16.2923 | -35.1355 | -27.0219
16.2984 | -35.136 | -27.03
16.3045 | -35.1366 | -27.0379
16.3106 | -35.1372 | -27.0461
16.3167 | -35.1378 | -27.0544
16.3228 | -35.1384 | -27.0622
16.3289 | -35.1389 | -27.0703
16.335 | -35.1395 | -27.0783
16.3411 | -35.14 | -27.0862
16.3472 | -35.1406 | -27.0944
16.3533 | -35.1412 | -27.1024
16.3594 | -35.1418 | -27.1101
16.3655 | -35.1423 | -27.118
16.3716 | -35.1429 | -27.1258
16.3777 | -35.1435 | -27.1337
16.3838 | -35.144 | -27.1417
16.3899 | -35.1446 | -27.1496
16.396 | -35.1452 | -27.1577
16.4021 | -35.1458 | -27.1658
16.4083 | -35.1463 | -27.1735
16.4144 | -35.1469 | -27.1817
16.4205 | -35.1475 | -27.1897
16.4266 | -35.148 | -27.1976
16.4327 | -35.1486 | -27.2055
16.4388 | -35.1492 | -27.2135
16.4449 | -35.1497 | -27.2214
16.451 | -35.1503 | -27.2294
16.4571 | -35.1509 | -27.2375
16.4632 | -35.1514 | -27.2452
16.4693 | -35.152 | -27.2531
16.4754 | -35.1526 | -27.2613
16.4815 | -35.1531 | -27.2691
16.4876 | -35.1537 | -27.2772
16.4937 | -35.1543 | -27.2849
16.4998 | -35.1549 | -27.2928
16.5059 | -35.1554 | -27.3007
16.512 | -35.156 | -27.3087
16.5181 | -35.1566 | -27.3167
16.5242 | -35.1572 | -27.3248
16.5303 | -35.1578 | -27.3325
16.5364 | -35.1583 | -27.3404
16.5425 | -35.1589 | -27.3486
16.5487 | -35.1595 | -27.3563
16.5548 | -35.16 | -27.3643
16.5609 | -35.1606 | -27.3723
16.567 | -35.1612 | -27.38
16.5731 | -35.1618 | -27.3878
16.5792 | -35.1624 | -27.3958
16.5853 | -35.1629 | -27.4037
16.5914 | -35.1635 | -27.4117
16.5975 | -35.1641 | -27.4198
16.6036 | -35.1647 | -27.4276
16.6097 | -35.1652 | -27.4356
16.6158 | -35.1658 | -27.4436
16.6219 | -35.1664 | -27.4516
16.628 | -35.1669 | -27.4594
16.6341 | -35.1675 | -27.4673
16.6402 | -35.168 | -27.475
16.6463 | -35.1686 | -27.4831
16.6524 | -35.1692 | -27.491
16.6585 | -35.1697 | -27.4989
16.6646 | -35.1703 | -27.5067
16.6707 | -35.1709 | -27.5143
16.6768 | -35.1715 | -27.5222
16.6829 | -35.1721 | -27.53
16.689 | -35.1727 | -27.5378
16.6952 | -35.1732 | -27.5455
16.7013 | -35.1738 | -27.5534
16.7074 | -35.1744 | -27.561
16.7135 | -35.1749 | -27.5688
16.7196 | -35.1755 | -27.5768
16.7257 | -35.1761 | -27.5847
16.7318 | -35.1767 | -27.5925
16.7379 | -35.1772 | -27.6002
16.744 | -35.1778 | -27.6078
16.7501 | -35.1784 | -27.6157
16.7562 | -35.1789 | -27.6236
16.7623 | -35.1795 | -27.6314
16.7684 | -35.1801 | -27.6392
16.7745 | -35.1806 | -27.6472
16.7806 | -35.1812 | -27.6548
16.7867 | -35.1818 | -27.6626
16.7928 | -35.1824 | -27.6704
16.7989 | -35.1829 | -27.6782
16.805 | -35.1835 | -27.6862
16.8111 | -35.1841 | -27.6938
16.8172 | -35.1846 | -27.7018
16.8233 | -35.1852 | -27.7097
16.8294 | -35.1858 | -27.7176
16.8356 | -35.1864 | -27.7255
16.8417 | -35.187 | -27.7335
16.8478 | -35.1876 | -27.7412
16.8539 | -35.1881 | -27.749
16.86 | -35.1887 | -27.7569
16.8661 | -35.1893 | -27.7646
16.8722 | -35.1898 | -27.7728
16.8783 | -35.1904 | -27.7808
16.8844 | -35.191 | -27.7884
16.8905 | -35.1916 | -27.7962
16.8966 | -35.1921 | -27.804
16.9027 | -35.1927 | -27.8117
16.9088 | -35.1933 | -27.8196
16.9149 | -35.1939 | -27.8275
16.921 | -35.1945 | -27.8351
16.9271 | -35.195 | -27.843
16.9332 | -35.1956 | -27.8508
16.9393 | -35.1962 | -27.8587
16.9454 | -35.1967 | -27.8664
16.9515 | -35.1973 | -27.8742
16.9576 | -35.1979 | -27.8817
16.9637 | -35.1985 | -27.8895
16.9698 | -35.1991 | -27.8972
16.9759 | -35.1997 | -27.9051
16.9821 | -35.2003 | -27.9131
16.9882 | -35.2009 | -27.9206
16.9943 | -35.2014 | -27.9284
17.0004 | -35.202 | -27.9362
17.0065 | -35.2026 | -27.9441
17.0126 | -35.2032 | -27.952
17.0187 | -35.2037 | -27.96
17.0248 | -35.2043 | -27.9676
17.0309 | -35.2049 | -27.9754
17.037 | -35.2054 | -27.9831
17.0431 | -35.206 | -27.9909
17.0492 | -35.2066 | -27.9987
17.0553 | -35.2071 | -28.0064
17.0614 | -35.2077 | -28.0139
17.0675 | -35.2083 | -28.0216
17.0736 | -35.2089 | -28.0294
17.0797 | -35.2095 | -28.037
17.0858 | -35.21 | -28.0448
17.0919 | -35.2106 | -28.0525
17.098 | -35.2112 | -28.0599
17.1041 | -35.2118 | -28.0677
17.1102 | -35.2123 | -28.0755
17.1163 | -35.2129 | -28.0832
17.1225 | -35.2135 | -28.0908
17.1286 | -35.2141 | -28.0982
17.1347 | -35.2147 | -28.106
17.1408 | -35.2153 | -28.1139
17.1469 | -35.2159 | -28.1215
17.153 | -35.2164 | -28.1293
17.1591 | -35.217 | -28.137
17.1652 | -35.2176 | -28.1446
17.1713 | -35.2182 | -28.1523
17.1774 | -35.2187 | -28.1602
17.1835 | -35.2193 | -28.1679
17.1896 | -35.2199 | -28.1756
17.1957 | -35.2205 | -28.1835
17.2018 | -35.221 | -28.1909
17.2079 | -35.2216 | -28.1987
17.214 | -35.2222 | -28.2066
17.2201 | -35.2228 | -28.2143
17.2262 | -35.2233 | -28.2221
17.2323 | -35.2239 | -28.2299
17.2384 | -35.2245 | -28.2375
17.2445 | -35.2251 | -28.2452
17.2506 | -35.2256 | -28.2528
17.2567 | -35.2262 | -28.2606
17.2628 | -35.2268 | -28.2683
17.269 | -35.2274 | -28.2759
17.2751 | -35.2279 | -28.2836
17.2812 | -35.2285 | -28.2915
17.2873 | -35.2291 | -28.2992
17.2934 | -35.2297 | -28.307
17.2995 | -35.2302 | -28.3148
17.3056 | -35.2308 | -28.3222
17.3117 | -35.2314 | -28.3301
17.3178 | -35.232 | -28.3377
17.3239 | -35.2326 | -28.3454
17.33 | -35.2331 | -28.353
17.3361 | -35.2337 | -28.3606
17.3422 | -35.2342 | -28.368
17.3483 | -35.2348 | -28.3757
17.3544 | -35.2354 | -28.3836
17.3605 | -35.236 | -28.3913
17.3666 | -35.2366 | -28.3991
17.3727 | -35.2371 | -28.4069
17.3788 | -35.2377 | -28.4143
17.3849 | -35.2383 | -28.422
17.391 | -35.2388 | -28.4296
17.3971 | -35.2394 | -28.4372
17.4032 | -35.24 | -28.4449
17.4094 | -35.2406 | -28.4524
17.4155 | -35.2411 | -28.46
17.4216 | -35.2417 | -28.4676
17.4277 | -35.2423 | -28.4753
17.4338 | -35.2428 | -28.483
17.4399 | -35.2434 | -28.4907
17.446 | -35.244 | -28.498
17.4521 | -35.2446 | -28.5056
17.4582 | -35.2451 | -28.513
17.4643 | -35.2457 | -28.5207
17.4704 | -35.2463 | -28.5284
17.4765 | -35.2469 | -28.5361
17.4826 | -35.2474 | -28.5436
17.4887 | -35.248 | -28.5514
17.4948 | -35.2486 | -28.5591
17.5009 | -35.2491 | -28.5667
17.507 | -35.2497 | -28.5745
17.5131 | -35.2503 | -28.5821
17.5192 | -35.2508 | -28.5897
17.5253 | -35.2514 | -28.5973
17.5314 | -35.252 | -28.6049
17.5375 | -35.2525 | -28.6126
17.5436 | -35.2531 | -28.6202
17.5497 | -35.2537 | -28.6278
17.5559 | -35.2542 | -28.6351
17.562 | -35.2548 | -28.6426
17.5681 | -35.2554 | -28.6501
17.5742 | -35.256 | -28.6578
17.5803 | -35.2566 | -28.6653
17.5864 | -35.2571 | -28.6729
17.5925 | -35.2577 | -28.6805
17.5986 | -35.2583 | -28.6882
17.6047 | -35.2589 | -28.6958
17.6108 | -35.2595 | -28.7034
17.6169 | -35.26 | -28.7109
17.623 | -35.2606 | -28.7183
17.6291 | -35.2612 | -28.7258
17.6352 | -35.2618 | -28.7334
17.6413 | -35.2623 | -28.7411
17.6474 | -35.2629 | -28.7488
17.6535 | -35.2635 | -28.7564
17.6596 | -35.264 | -28.7638
17.6657 | -35.2646 | -28.7715
17.6718 | -35.2652 | -28.7791
17.6779 | -35.2658 | -28.7866
17.684 | -35.2663 | -28.7941
17.6901 | -35.2669 | -28.8016
17.6963 | -35.2675 | -28.809
17.7024 | -35.268 | -28.8166
17.7085 | -35.2686 | -28.8241
17.7146 | -35.2692 | -28.8316
17.7207 | -35.2698 | -28.8391
17.7268 | -35.2704 | -28.8465
17.7329 | -35.2709 | -28.8541
17.739 | -35.2715 | -28.8617
17.7451 | -35.2721 | -28.8691
17.7512 | -35.2726 | -28.8766
17.7573 | -35.2732 | -28.8842
17.7634 | -35.2738 | -28.8914
17.7695 | -35.2744 | -28.899
17.7756 | -35.275 | -28.9065
17.7817 | -35.2755 | -28.914
17.7878 | -35.2761 | -28.9217
17.7939 | -35.2767 | -28.9292
17.8 | -35.2772 | -28.9367
17.8061 | -35.2778 | -28.9443
17.8122 | -35.2784 | -28.9518
17.8183 | -35.279 | -28.9595
17.8244 | -35.2795 | -28.967
17.8305 | -35.2801 | -28.9746
17.8366 | -35.2807 | -28.9819
17.8428 | -35.2813 | -28.9895
17.8489 | -35.2818 | -28.9969
17.855 | -35.2824 | -29.0044
17.8611 | -35.283 | -29.0119
17.8672 | -35.2836 | -29.0193
17.8733 | -35.2841 | -29.0268
17.8794 | -35.2847 | -29.0343
17.8855 | -35.2853 | -29.0418
17.8916 | -35.2858 | -29.0493
17.8977 | -35.2864 | -29.0568
17.9038 | -35.287 | -29.064
17.9099 | -35.2875 | -29.0715
17.916 | -35.2881 | -29.0789
17.9221 | -35.2887 | -29.0863
17.9282 | -35.2893 | -29.0938
17.9343 | -35.2899 | -29.1014
17.9404 | -35.2904 | -29.1087
17.9465 | -35.291 | -29.1162
17.9526 | -35.2916 | -29.1239
17.9587 | -35.2921 | -29.1315
17.9648 | -35.2927 | -29.139
17.9709 | -35.2932 | -29.1465
17.977 | -35.2938 | -29.154
17.9832 | -35.2944 | -29.1615
17.9893 | -35.2949 | -29.169
17.9954 | -35.2955 | -29.1766
18.0015 | -35.296 | -29.1842
18.0076 | -35.2966 | -29.1914
18.0137 | -35.2972 | -29.1989
18.0198 | -35.2977 | -29.2064
18.0259 | -35.2983 | -29.2138
18.032 | -35.2989 | -29.2213
18.0381 | -35.2994 | -29.2289
18.0442 | -35.3 | -29.2363
18.0503 | -35.3006 | -29.2438
18.0564 | -35.3011 | -29.2513
18.0625 | -35.3017 | -29.259
18.0686 | -35.3023 | -29.2665
18.0747 | -35.3028 | -29.2741
18.0808 | -35.3034 | -29.2812
18.0869 | -35.304 | -29.2887
18.093 | -35.3046 | -29.2962
18.0991 | -35.3051 | -29.3036
18.1052 | -35.3056 | -29.3113
18.1113 | -35.3062 | -29.3189
18.1174 | -35.3068 | -29.3261
18.1236 | -35.3074 | -29.3333
18.1297 | -35.3079 | -29.3408
18.1358 | -35.3085 | -29.3483
18.1419 | -35.3091 | -29.3557
18.148 | -35.3097 | -29.3631
18.1541 | -35.3102 | -29.3703
18.1602 | -35.3108 | -29.3778
18.1663 | -35.3114 | -29.3852
18.1724 | -35.312 | -29.3927
18.1785 | -35.3125 | -29.4002
18.1846 | -35.313 | -29.4076
18.1907 | -35.3136 | -29.4151
18.1968 | -35.3142 | -29.4227
18.2029 | -35.3147 | -29.4302
18.209 | -35.3153 | -29.4378
18.2151 | -35.3159 | -29.4453
18.2212 | -35.3165 | -29.4526
18.2273 | -35.3171 | -29.46
18.2334 | -35.3176 | -29.4675
18.2395 | -35.3182 | -29.4749
18.2456 | -35.3187 | -29.4826
18.2517 | -35.3193 | -29.49
18.2578 | -35.3199 | -29.4972
18.2639 | -35.3204 | -29.5048
18.2701 | -35.321 | -29.5122
18.2762 | -35.3216 | -29.5196
18.2823 | -35.3221 | -29.5271
18.2884 | -35.3227 | -29.5345
18.2945 | -35.3233 | -29.5416
18.3006 | -35.3239 | -29.549
18.3067 | -35.3245 | -29.5564
18.3128 | -35.325 | -29.5641
18.3189 | -35.3256 | -29.5714
18.325 | -35.3262 | -29.5788
18.3311 | -35.3267 | -29.5862
18.3372 | -35.3273 | -29.5936
18.3433 | -35.3278 | -29.601
18.3494 | -35.3284 | -29.6085
18.3555 | -35.329 | -29.6159
18.3616 | -35.3296 | -29.6232
18.3677 | -35.3301 | -29.6307
18.3738 | -35.3307 | -29.6381
18.3799 | -35.3313 | -29.6456
18.386 | -35.3318 | -29.6531
18.3921 | -35.3324 | -29.6604
18.3982 | -35.3329 | -29.6675
18.4043 | -35.3335 | -29.675
18.4105 | -35.334 | -29.6825
18.4166 | -35.3346 | -29.6898
18.4227 | -35.3352 | -29.6973
18.4288 | -35.3357 | -29.7047
18.4349 | -35.3363 | -29.712
18.441 | -35.3369 | -29.7194
18.4471 | -35.3374 | -29.7267
18.4532 | -35.338 | -29.7341
18.4593 | -35.3386 | -29.7414
18.4654 | -35.3391 | -29.7486
18.4715 | -35.3397 | -29.7561
18.4776 | -35.3402 | -29.7636
18.4837 | -35.3408 | -29.7711
18.4898 | -35.3414 | -29.7786
18.4959 | -35.3419 | -29.7862
18.502 | -35.3425 | -29.7935
18.5081 | -35.343 | -29.8009
18.5142 | -35.3435 | -29.8086
18.5203 | -35.3441 | -29.816
18.5264 | -35.3447 | -29.8234
18.5325 | -35.3453 | -29.831
18.5386 | -35.3458 | -29.8384
18.5447 | -35.3464 | -29.8459
18.5508 | -35.3469 | -29.8531
18.557 | -35.3475 | -29.8608
18.5631 | -35.3481 | -29.8683
18.5692 | -35.3487 | -29.8756
18.5753 | -35.3493 | -29.883
18.5814 | -35.3498 | -29.8903
18.5875 | -35.3504 | -29.8977
18.5936 | -35.3509 | -29.9051
18.5997 | -35.3515 | -29.9126
18.6058 | -35.352 | -29.9199
18.6119 | -35.3526 | -29.9273
18.618 | -35.3532 | -29.9348
18.6241 | -35.3537 | -29.9423
18.6302 | -35.3543 | -29.9499
18.6363 | -35.3548 | -29.9574
18.6424 | -35.3554 | -29.9648
18.6485 | -35.356 | -29.9722
18.6546 | -35.3566 | -29.9797
18.6607 | -35.3571 | -29.9871
18.6668 | -35.3577 | -29.9945
18.6729 | -35.3582 | -30.0021
18.679 | -35.3588 | -30.0093
18.6851 | -35.3594 | -30.0166
18.6912 | -35.3599 | -30.024
18.6974 | -35.3605 | -30.0314
18.7035 | -35.3611 | -30.039
18.7096 | -35.3617 | -30.0465
18.7157 | -35.3622 | -30.0538
18.7218 | -35.3628 | -30.0613
18.7279 | -35.3634 | -30.0687
18.734 | -35.364 | -30.0762
18.7401 | -35.3645 | -30.0836
18.7462 | -35.3651 | -30.0908
18.7523 | -35.3657 | -30.0981
18.7584 | -35.3663 | -30.1056
18.7645 | -35.3669 | -30.1131
18.7706 | -35.3674 | -30.1205
18.7767 | -35.368 | -30.1278
18.7828 | -35.3685 | -30.135
18.7889 | -35.3691 | -30.1425
18.795 | -35.3697 | -30.1499
18.8011 | -35.3703 | -30.1574
18.8072 | -35.3708 | -30.1649
18.8133 | -35.3714 | -30.1722
18.8194 | -35.372 | -30.1792
18.8255 | -35.3725 | -30.1866
18.8316 | -35.3731 | -30.1941
18.8377 | -35.3736 | -30.2014
18.8439 | -35.3742 | -30.2087
18.85 | -35.3747 | -30.216
18.8561 | -35.3753 | -30.2232
18.8622 | -35.3759 | -30.2307
18.8683 | -35.3764 | -30.2378
18.8744 | -35.377 | -30.2451
18.8805 | -35.3775 | -30.2526
18.8866 | -35.3781 | -30.2602
18.8927 | -35.3787 | -30.2674
18.8988 | -35.3793 | -30.2748
18.9049 | -35.3798 | -30.2821
18.911 | -35.3804 | -30.2893
18.9171 | -35.381 | -30.2965
18.9232 | -35.3816 | -30.3038
18.9293 | -35.3822 | -30.3112
18.9354 | -35.3827 | -30.3186
18.9415 | -35.3834 | -30.3259
18.9476 | -35.3839 | -30.3331
18.9537 | -35.3845 | -30.3405
18.9598 | -35.385 | -30.3477
18.9659 | -35.3856 | -30.3551
18.972 | -35.3862 | -30.3625
18.9781 | -35.3868 | -30.37
18.9843 | -35.3873 | -30.3772
18.9904 | -35.3879 | -30.3846
18.9965 | -35.3885 | -30.3918
19.0026 | -35.389 | -30.3992
19.0087 | -35.3896 | -30.4066
19.0148 | -35.3902 | -30.414
19.0209 | -35.3907 | -30.4213
19.027 | -35.3913 | -30.4287
19.0331 | -35.3918 | -30.4359
19.0392 | -35.3924 | -30.4431
19.0453 | -35.393 | -30.4506
19.0514 | -35.3935 | -30.4577
19.0575 | -35.3941 | -30.4652
19.0636 | -35.3947 | -30.4724
19.0697 | -35.3953 | -30.4797
19.0758 | -35.3958 | -30.4871
19.0819 | -35.3964 | -30.4945
19.088 | -35.397 | -30.5015
19.0941 | -35.3975 | -30.5089
19.1002 | -35.398 | -30.5162
19.1063 | -35.3986 | -30.5237
19.1124 | -35.3991 | -30.531
19.1185 | -35.3997 | -30.5384
19.1246 | -35.4003 | -30.5457
19.1308 | -35.4009 | -30.5532
19.1369 | -35.4014 | -30.5605
19.143 | -35.402 | -30.5677
19.1491 | -35.4026 | -30.5751
19.1552 | -35.4031 | -30.5824
19.1613 | -35.4037 | -30.5898
19.1674 | -35.4042 | -30.597
19.1735 | -35.4048 | -30.6044
19.1796 | -35.4053 | -30.6116
19.1857 | -35.4059 | -30.619
19.1918 | -35.4065 | -30.6264
19.1979 | -35.407 | -30.6336
19.204 | -35.4076 | -30.6409
19.2101 | -35.4082 | -30.6481
19.2162 | -35.4088 | -30.6553
19.2223 | -35.4094 | -30.6629
19.2284 | -35.4099 | -30.6701
19.2345 | -35.4104 | -30.6775
19.2406 | -35.411 | -30.6847
19.2467 | -35.4115 | -30.6921
19.2528 | -35.4121 | -30.6996
19.2589 | -35.4127 | -30.7067
19.265 | -35.4133 | -30.714
19.2712 | -35.4138 | -30.7213
19.2773 | -35.4144 | -30.7285
19.2834 | -35.415 | -30.7356
19.2895 | -35.4156 | -30.743
19.2956 | -35.4161 | -30.7501
19.3017 | -35.4167 | -30.7573
19.3078 | -35.4172 | -30.7647
19.3139 | -35.4178 | -30.772
19.32 | -35.4184 | -30.779
19.3261 | -35.4189 | -30.7862
19.3322 | -35.4195 | -30.7932
19.3383 | -35.4201 | -30.8004
19.3444 | -35.4207 | -30.8076
19.3505 | -35.4212 | -30.815
19.3566 | -35.4218 | -30.8222
19.3627 | -35.4223 | -30.8294
19.3688 | -35.423 | -30.8363
19.3749 | -35.4235 | -30.8437
19.381 | -35.4241 | -30.8509
19.3871 | -35.4247 | -30.8579
19.3932 | -35.4252 | -30.865
19.3993 | -35.4258 | -30.8725
19.4054 | -35.4264 | -30.8797
19.4115 | -35.4269 | -30.8871
19.4177 | -35.4275 | -30.8942
19.4238 | -35.428 | -30.9015
19.4299 | -35.4286 | -30.9087
19.436 | -35.4292 | -30.916
19.4421 | -35.4298 | -30.9232
19.4482 | -35.4303 | -30.9306
19.4543 | -35.4309 | -30.9376
19.4604 | -35.4314 | -30.945
19.4665 | -35.432 | -30.9522
19.4726 | -35.4326 | -30.9593
19.4787 | -35.4331 | -30.9665
19.4848 | -35.4337 | -30.9738
19.4909 | -35.4342 | -30.9809
19.497 | -35.4348 | -30.9882
19.5031 | -35.4353 | -30.9953
19.5092 | -35.4359 | -31.0027
19.5153 | -35.4364 | -31.0098
19.5214 | -35.437 | -31.0168
19.5275 | -35.4375 | -31.0241
19.5336 | -35.4381 | -31.0314
19.5397 | -35.4386 | -31.0387
19.5458 | -35.4392 | -31.0458
19.5519 | -35.4397 | -31.0531
19.5581 | -35.4403 | -31.0601
19.5642 | -35.4409 | -31.0673
19.5703 | -35.4414 | -31.0745
19.5764 | -35.442 | -31.0817
19.5825 | -35.4425 | -31.0889
19.5886 | -35.4431 | -31.0963
19.5947 | -35.4437 | -31.1032
19.6008 | -35.4442 | -31.1104
19.6069 | -35.4448 | -31.1176
19.613 | -35.4454 | -31.125
19.6191 | -35.4459 | -31.1322
19.6252 | -35.4464 | -31.1394
19.6313 | -35.447 | -31.1464
19.6374 | -35.4475 | -31.1537
19.6435 | -35.4481 | -31.1608
19.6496 | -35.4487 | -31.168
19.6557 | -35.4492 | -31.1754
19.6618 | -35.4498 | -31.1824
19.6679 | -35.4503 | -31.1897
19.674 | -35.4509 | -31.1969
19.6801 | -35.4515 | -31.2039
19.6862 | -35.4521 | -31.2111
19.6923 | -35.4526 | -31.218
19.6984 | -35.4532 | -31.2253
19.7046 | -35.4538 | -31.2326
19.7107 | -35.4543 | -31.2398
19.7168 | -35.4549 | -31.247
19.7229 | -35.4554 | -31.254
19.729 | -35.456 | -31.2612
19.7351 | -35.4565 | -31.2681
19.7412 | -35.4571 | -31.2756
19.7473 | -35.4577 | -31.2827
19.7534 | -35.4582 | -31.2899
19.7595 | -35.4588 | -31.297
19.7656 | -35.4593 | -31.3042
19.7717 | -35.4599 | -31.3111
19.7778 | -35.4605 | -31.3182
19.7839 | -35.461 | -31.3255
19.79 | -35.4616 | -31.3325
19.7961 | -35.4621 | -31.3396
19.8022 | -35.4627 | -31.3466
19.8083 | -35.4633 | -31.3538
19.8144 | -35.4639 | -31.3608
19.8205 | -35.4644 | -31.3679
19.8266 | -35.465 | -31.375
19.8327 | -35.4655 | -31.3822
19.8388 | -35.4661 | -31.3892
19.845 | -35.4667 | -31.3963
19.8511 | -35.4673 | -31.4037
19.8572 | -35.4678 | -31.4107
19.8633 | -35.4684 | -31.4178
19.8694 | -35.4689 | -31.4254
19.8755 | -35.4695 | -31.4325
19.8816 | -35.47 | -31.4396
19.8877 | -35.4706 | -31.4468
19.8938 | -35.4711 | -31.4541
19.8999 | -35.4717 | -31.4614
19.906 | -35.4722 | -31.4685
19.9121 | -35.4728 | -31.4754
19.9182 | -35.4733 | -31.4826
19.9243 | -35.4738 | -31.4896
19.9304 | -35.4744 | -31.4967
19.9365 | -35.475 | -31.5039
19.9426 | -35.4756 | -31.5108
19.9487 | -35.4761 | -31.5179
19.9548 | -35.4767 | -31.5252
19.9609 | -35.4773 | -31.5323
19.967 | -35.4778 | -31.5393
19.9731 | -35.4784 | -31.5464
19.9792 | -35.479 | -31.5534
19.9853 | -35.4795 | -31.5606
19.9915 | -35.4801 | -31.5678
19.9976 | -35.4807 | -31.5748
20.0037 | -35.4812 | -31.5821
20.0098 | -35.4818 | -31.5894
20.0159 | -35.4823 | -31.5962
20.022 | -35.4829 | -31.6032
20.0281 | -35.4834 | -31.6101
20.0342 | -35.484 | -31.6171
20.0403 | -35.4846 | -31.6242
20.0464 | -35.4851 | -31.6314
20.0525 | -35.4857 | -31.6382
20.0586 | -35.4863 | -31.6452
20.0647 | -35.4869 | -31.6524
20.0708 | -35.4874 | -31.6597
20.0769 | -35.488 | -31.667
20.083 | -35.4886 | -31.6739
20.0891 | -35.4892 | -31.6808
20.0952 | -35.4897 | -31.6879
20.1013 | -35.4903 | -31.6949
20.1074 | -35.4908 | -31.7021
20.1135 | -35.4914 | -31.7096
20.1196 | -35.492 | -31.7162
20.1257 | -35.4925 | -31.7233
20.1319 | -35.4931 | -31.7303
20.138 | -35.4936 | -31.7374
20.1441 | -35.4942 | -31.7446
20.1502 | -35.4947 | -31.7517
20.1563 | -35.4953 | -31.7585
20.1624 | -35.4959 | -31.7657
20.1685 | -35.4964 | -31.7728
20.1746 | -35.497 | -31.78
20.1807 | -35.4975 | -31.7871
20.1868 | -35.4981 | -31.7942
20.1929 | -35.4986 | -31.801
20.199 | -35.4992 | -31.808
20.2051 | -35.4998 | -31.815
20.2112 | -35.5003 | -31.8222
20.2173 | -35.5009 | -31.8293
20.2234 | -35.5014 | -31.8362
20.2295 | -35.502 | -31.8431
20.2356 | -35.5025 | -31.8502
20.2417 | -35.5031 | -31.8573
20.2478 | -35.5036 | -31.8644
20.2539 | -35.5041 | -31.8714
20.26 | -35.5047 | -31.878
20.2661 | -35.5053 | -31.885
20.2723 | -35.5059 | -31.8921
20.2784 | -35.5065 | -31.8991
20.2845 | -35.5071 | -31.906
20.2906 | -35.5076 | -31.9131
20.2967 | -35.5082 | -31.9199
20.3028 | -35.5088 | -31.9268
20.3089 | -35.5093 | -31.9339
20.315 | -35.5099 | -31.9411
20.3211 | -35.5104 | -31.948
20.3272 | -35.511 | -31.9551
20.3333 | -35.5115 | -31.9619
20.3394 | -35.5121 | -31.9689
20.3455 | -35.5126 | -31.9759
20.3516 | -35.5132 | -31.9828
20.3577 | -35.5137 | -31.9899
20.3638 | -35.5143 | -31.9969
20.3699 | -35.5149 | -32.0036
20.376 | -35.5155 | -32.0105
20.3821 | -35.5161 | -32.0176
20.3882 | -35.5166 | -32.0246
20.3943 | -35.5172 | -32.0315
20.4004 | -35.5178 | -32.0383
20.4065 | -35.5183 | -32.0452
20.4126 | -35.5189 | -32.0522
20.4188 | -35.5195 | -32.0594
20.4249 | -35.52 | -32.0662
20.431 | -35.5206 | -32.073
20.4371 | -35.5211 | -32.0801
20.4432 | -35.5217 | -32.0871
20.4493 | -35.5222 | -32.0942
20.4554 | -35.5228 | -32.1012
20.4615 | -35.5234 | -32.1083
20.4676 | -35.5239 | -32.1154
20.4737 | -35.5245 | -32.1222
20.4798 | -35.5251 | -32.1291
20.4859 | -35.5257 | -32.1361
20.492 | -35.5263 | -32.1432
20.4981 | -35.5269 | -32.1501
20.5042 | -35.5274 | -32.1571
20.5103 | -35.528 | -32.1639
20.5164 | -35.5285 | -32.1709
20.5225 | -35.5291 | -32.1778
20.5286 | -35.5296 | -32.185
20.5347 | -35.5302 | -32.1918
20.5408 | -35.5308 | -32.1986
20.5469 | -35.5313 | -32.2056
20.553 | -35.5319 | -32.2127
20.5592 | -35.5325 | -32.2197
20.5653 | -35.533 | -32.2266
20.5714 | -35.5336 | -32.2335
20.5775 | -35.5342 | -32.2403
20.5836 | -35.5348 | -32.2472
20.5897 | -35.5353 | -32.2542
20.5958 | -35.5359 | -32.2613
20.6019 | -35.5364 | -32.2685
20.608 | -35.537 | -32.2755
20.6141 | -35.5376 | -32.2825
20.6202 | -35.5381 | -32.2894
20.6263 | -35.5387 | -32.2966
20.6324 | -35.5392 | -32.3035
20.6385 | -35.5398 | -32.3107
20.6446 | -35.5404 | -32.3177
20.6507 | -35.5409 | -32.3245
20.6568 | -35.5415 | -32.3316
20.6629 | -35.542 | -32.3387
20.669 | -35.5426 | -32.3456
20.6751 | -35.5432 | -32.3527
20.6812 | -35.5438 | -32.3591
20.6873 | -35.5443 | -32.3661
20.6934 | -35.5448 | -32.3731
20.6995 | -35.5454 | -32.3801
20.7057 | -35.546 | -32.3874
20.7118 | -35.5466 | -32.3944
20.7179 | -35.5471 | -32.4008
20.724 | -35.5477 | -32.4079
20.7301 | -35.5483 | -32.4145
20.7362 | -35.5488 | -32.4213
20.7423 | -35.5494 | -32.4284
20.7484 | -35.55 | -32.4354
20.7545 | -35.5505 | -32.4421
20.7606 | -35.5511 | -32.4493
20.7667 | -35.5517 | -32.456
20.7728 | -35.5522 | -32.4631
20.7789 | -35.5528 | -32.4702
20.785 | -35.5534 | -32.4772
20.7911 | -35.5539 | -32.484
20.7972 | -35.5544 | -32.4912
20.8033 | -35.555 | -32.4982
20.8094 | -35.5556 | -32.505
20.8155 | -35.5561 | -32.5119
20.8216 | -35.5567 | -32.5189
20.8277 | -35.5573 | -32.5255
20.8338 | -35.5579 | -32.5324
20.8399 | -35.5585 | -32.5394
20.8461 | -35.559 | -32.5466
20.8522 | -35.5596 | -32.5535
20.8583 | -35.5603 | -32.5598
20.8644 | -35.5608 | -32.567
20.8705 | -35.5614 | -32.5739
20.8766 | -35.5619 | -32.5808
20.8827 | -35.5625 | -32.5877
20.8888 | -35.5631 | -32.5945
20.8949 | -35.5637 | -32.6014
20.901 | -35.5643 | -32.6084
20.9071 | -35.5648 | -32.6152
20.9132 | -35.5654 | -32.6219
20.9193 | -35.5659 | -32.6288
20.9254 | -35.5665 | -32.6356
20.9315 | -35.5671 | -32.6424
20.9376 | -35.5676 | -32.6494
20.9437 | -35.5682 | -32.656
20.9498 | -35.5689 | -32.6629
20.9559 | -35.5695 | -32.6695
20.962 | -35.57 | -32.6763
20.9681 | -35.5706 | -32.6832
20.9742 | -35.5711 | -32.6904
20.9803 | -35.5717 | -32.6974
20.9864 | -35.5723 | -32.7043
20.9926 | -35.5728 | -32.7112
20.9987 | -35.5734 | -32.7179
21.0048 | -35.5739 | -32.7245
21.0109 | -35.5745 | -32.7313
21.017 | -35.5751 | -32.738
21.0231 | -35.5757 | -32.7447
21.0292 | -35.5763 | -32.7517
21.0353 | -35.5769 | -32.7579
21.0414 | -35.5775 | -32.7649
21.0475 | -35.578 | -32.7717
21.0536 | -35.5786 | -32.7782
21.0597 | -35.5791 | -32.7852
21.0658 | -35.5797 | -32.7921
21.0719 | -35.5802 | -32.7988
21.078 | -35.5808 | -32.8058
21.0841 | -35.5814 | -32.8126
21.0902 | -35.5819 | -32.8194
21.0963 | -35.5825 | -32.8261
21.1024 | -35.5831 | -32.8332
21.1085 | -35.5838 | -32.8399
21.1146 | -35.5843 | -32.8463
21.1207 | -35.5849 | -32.8531
21.1268 | -35.5854 | -32.8599
21.133 | -35.5859 | -32.8668
21.1391 | -35.5865 | -32.8733
21.1452 | -35.587 | -32.8801
21.1513 | -35.5876 | -32.8867
21.1574 | -35.5881 | -32.8937
21.1635 | -35.5887 | -32.9008
21.1696 | -35.5892 | -32.9073
21.1757 | -35.5898 | -32.9137
21.1818 | -35.5904 | -32.9206
21.1879 | -35.5909 | -32.9278
21.194 | -35.5915 | -32.9347
21.2001 | -35.5921 | -32.941
21.2062 | -35.5926 | -32.9477
21.2123 | -35.5932 | -32.9541
21.2184 | -35.5938 | -32.9608
21.2245 | -35.5944 | -32.9676
21.2306 | -35.595 | -32.9742
21.2367 | -35.5956 | -32.9811
21.2428 | -35.5961 | -32.9875
21.2489 | -35.5967 | -32.9934
21.255 | -35.5973 | -33.0002
21.2611 | -35.5978 | -33.007
21.2672 | -35.5983 | -33.0135
21.2733 | -35.5989 | -33.0205
21.2795 | -35.5995 | -33.0269
21.2856 | -35.6 | -33.0335
21.2917 | -35.6006 | -33.0402
21.2978 | -35.6011 | -33.0468
21.3039 | -35.6017 | -33.0534
21.31 | -35.6023 | -33.0603
21.3161 | -35.6029 | -33.0668
21.3222 | -35.6035 | -33.0734
21.3283 | -35.604 | -33.0802
21.3344 | -35.6046 | -33.0867
21.3405 | -35.6052 | -33.0936
21.3466 | -35.6058 | -33.1004
21.3527 | -35.6063 | -33.1068
21.3588 | -35.6069 | -33.1134
21.3649 | -35.6075 | -33.1201
21.371 | -35.608 | -33.127
21.3771 | -35.6086 | -33.1335
21.3832 | -35.6092 | -33.1404
21.3893 | -35.6098 | -33.1471
21.3954 | -35.6103 | -33.1538
21.4015 | -35.6109 | -33.1604
21.4076 | -35.6115 | -33.1672
21.4137 | -35.612 | -33.1738
21.4199 | -35.6126 | -33.1805
21.426 | -35.6131 | -33.1871
21.4321 | -35.6137 | -33.1938
21.4382 | -35.6143 | -33.2004
21.4443 | -35.6148 | -33.2069
21.4504 | -35.6154 | -33.2137
21.4565 | -35.616 | -33.2201
21.4626 | -35.6165 | -33.2264
21.4687 | -35.6171 | -33.233
21.4748 | -35.6176 | -33.2397
21.4809 | -35.6182 | -33.2466
21.487 | -35.6187 | -33.2532
21.4931 | -35.6193 | -33.2599
21.4992 | -35.6198 | -33.2668
21.5053 | -35.6204 | -33.2735
21.5114 | -35.621 | -33.2803
21.5175 | -35.6215 | -33.2869
21.5236 | -35.6221 | -33.2935
21.5297 | -35.6227 | -33.3001
21.5358 | -35.6232 | -33.3067
21.5419 | -35.6238 | -33.3134
21.548 | -35.6243 | -33.3201
21.5541 | -35.6249 | -33.3268
21.5602 | -35.6254 | -33.3336
21.5664 | -35.626 | -33.34
21.5725 | -35.6266 | -33.3467
21.5786 | -35.6271 | -33.3534
21.5847 | -35.6277 | -33.3601
21.5908 | -35.6283 | -33.3667
21.5969 | -35.6288 | -33.373
21.603 | -35.6294 | -33.3797
21.6091 | -35.6299 | -33.3864
21.6152 | -35.6305 | -33.3933
21.6213 | -35.631 | -33.4
21.6274 | -35.6316 | -33.4066
21.6335 | -35.6321 | -33.4129
21.6396 | -35.6327 | -33.4198
21.6457 | -35.6333 | -33.4264
21.6518 | -35.6339 | -33.4326
21.6579 | -35.6344 | -33.4391
21.664 | -35.635 | -33.446
21.6701 | -35.6355 | -33.4524
21.6762 | -35.6361 | -33.4591
21.6823 | -35.6367 | -33.4659
21.6884 | -35.6372 | -33.4724
21.6945 | -35.6378 | -33.4789
21.7006 | -35.6383 | -33.4855
21.7068 | -35.6389 | -33.4917
21.7129 | -35.6395 | -33.4984
21.719 | -35.64 | -33.505
21.7251 | -35.6406 | -33.5116
21.7312 | -35.6412 | -33.5182
21.7373 | -35.6417 | -33.5246
21.7434 | -35.6423 | -33.5312
21.7495 | -35.6429 | -33.5379
21.7556 | -35.6434 | -33.5445
21.7617 | -35.644 | -33.5511
21.7678 | -35.6445 | -33.5576
21.7739 | -35.6451 | -33.5637
21.78 | -35.6457 | -33.5702
21.7861 | -35.6463 | -33.5771
21.7922 | -35.6468 | -33.5838
21.7983 | -35.6474 | -33.5905
21.8044 | -35.6479 | -33.5971
21.8105 | -35.6485 | -33.6033
21.8166 | -35.6491 | -33.61
21.8227 | -35.6497 | -33.6165
21.8288 | -35.6502 | -33.6235
21.8349 | -35.6508 | -33.63
21.841 | -35.6514 | -33.6366
21.8471 | -35.6519 | -33.6428
21.8533 | -35.6525 | -33.6494
21.8594 | -35.6531 | -33.6561
21.8655 | -35.6536 | -33.6627
21.8716 | -35.6542 | -33.6693
21.8777 | -35.6548 | -33.6758
21.8838 | -35.6553 | -33.6825
21.8899 | -35.6559 | -33.6891
21.896 | -35.6565 | -33.6955
21.9021 | -35.657 | -33.7021
21.9082 | -35.6576 | -33.7088
21.9143 | -35.6582 | -33.7151
21.9204 | -35.6588 | -33.7216
21.9265 | -35.6594 | -33.7284
21.9326 | -35.6599 | -33.7352
21.9387 | -35.6605 | -33.742
21.9448 | -35.6611 | -33.7484
21.9509 | -35.6616 | -33.7547
21.957 | -35.6622 | -33.7614
21.9631 | -35.6628 | -33.7677
21.9692 | -35.6633 | -33.7743
21.9753 | -35.6639 | -33.7808
21.9814 | -35.6645 | -33.7873
21.9875 | -35.665 | -33.7941
21.9937 | -35.6656 | -33.8004
21.9998 | -35.6662 | -33.807
22.0059 | -35.6667 | -33.8136
22.012 | -35.6673 | -33.8203
22.0181 | -35.6679 | -33.8263
22.0242 | -35.6685 | -33.8327
22.0303 | -35.669 | -33.8393
22.0364 | -35.6696 | -33.8462
22.0425 | -35.6701 | -33.8528
22.0486 | -35.6707 | -33.8592
22.0547 | -35.6712 | -33.8654
22.0608 | -35.6718 | -33.8718
22.0669 | -35.6723 | -33.8781
22.073 | -35.6728 | -33.8846
22.0791 | -35.6734 | -33.8909
22.0852 | -35.6739 | -33.8974
22.0913 | -35.6745 | -33.9039
22.0974 | -35.675 | -33.9103
22.1035 | -35.6756 | -33.917
22.1096 | -35.6761 | -33.9232
22.1157 | -35.6767 | -33.9299
22.1218 | -35.6773 | -33.936
22.1279 | -35.6779 | -33.9425
22.134 | -35.6785 | -33.9491
22.1402 | -35.6791 | -33.9556
22.1463 | -35.6796 | -33.9617
22.1524 | -35.6802 | -33.9682
22.1585 | -35.6807 | -33.9746
22.1646 | -35.6813 | -33.9809
22.1707 | -35.6818 | -33.9871
22.1768 | -35.6824 | -33.9936
22.1829 | -35.683 | -34.0002
22.189 | -35.6835 | -34.0067
22.1951 | -35.6841 | -34.013
22.2012 | -35.6846 | -34.0191
22.2073 | -35.6852 | -34.0254
22.2134 | -35.6857 | -34.0317
22.2195 | -35.6863 | -34.0384
22.2256 | -35.6869 | -34.0449
22.2317 | -35.6874 | -34.0512
22.2378 | -35.688 | -34.0577
22.2439 | -35.6885 | -34.0643
22.25 | -35.6891 | -34.0707
22.2561 | -35.6896 | -34.0771
22.2622 | -35.6902 | -34.0833
22.2683 | -35.6908 | -34.0896
22.2744 | -35.6913 | -34.0959
22.2806 | -35.6919 | -34.1026
22.2867 | -35.6924 | -34.1088
22.2928 | -35.693 | -34.1148
22.2989 | -35.6935 | -34.1213
22.305 | -35.6941 | -34.1274
22.3111 | -35.6946 | -34.1337
22.3172 | -35.6952 | -34.1402
22.3233 | -35.6958 | -34.1467
22.3294 | -35.6964 | -34.1532
22.3355 | -35.6969 | -34.1593
22.3416 | -35.6975 | -34.1657
22.3477 | -35.698 | -34.1723
22.3538 | -35.6986 | -34.1786
22.3599 | -35.6991 | -34.1851
22.366 | -35.6997 | -34.1914
22.3721 | -35.7002 | -34.1974
22.3782 | -35.7008 | -34.2038
22.3843 | -35.7013 | -34.2105
22.3904 | -35.7019 | -34.2167
22.3965 | -35.7024 | -34.2233
22.4026 | -35.7029 | -34.2301
22.4087 | -35.7035 | -34.2362
22.4148 | -35.704 | -34.2425
22.4209 | -35.7045 | -34.2489
22.4271 | -35.7051 | -34.2553
22.4332 | -35.7057 | -34.2616
22.4393 | -35.7062 | -34.268
22.4454 | -35.7068 | -34.2743
22.4515 | -35.7074 | -34.2808
22.4576 | -35.7079 | -34.2872
22.4637 | -35.7084 | -34.2936
22.4698 | -35.709 | -34.3
22.4759 | -35.7095 | -34.3059
22.482 | -35.7101 | -34.3127
22.4881 | -35.7106 | -34.3192
22.4942 | -35.7112 | -34.3256
22.5003 | -35.7117 | -34.3318
22.5064 | -35.7123 | -34.3382
22.5125 | -35.7129 | -34.3439
22.5186 | -35.7134 | -34.3504
22.5247 | -35.714 | -34.3568
22.5308 | -35.7145 | -34.3631
22.5369 | -35.7151 | -34.3696
22.543 | -35.7157 | -34.376
22.5491 | -35.7162 | -34.3821
22.5552 | -35.7168 | -34.3886
22.5613 | -35.7173 | -34.3953
22.5675 | -35.7179 | -34.4016
22.5736 | -35.7185 | -34.4081
22.5797 | -35.719 | -34.4145
22.5858 | -35.7196 | -34.4205
22.5919 | -35.7201 | -34.4268
22.598 | -35.7207 | -34.4332
22.6041 | -35.7213 | -34.4394
22.6102 | -35.7218 | -34.4458
22.6163 | -35.7224 | -34.4519
22.6224 | -35.723 | -34.4583
22.6285 | -35.7235 | -34.4646
22.6346 | -35.7241 | -34.4711
22.6407 | -35.7246 | -34.4774
22.6468 | -35.7252 | -34.4839
22.6529 | -35.7257 | -34.4897
22.659 | -35.7263 | -34.4959
22.6651 | -35.7269 | -34.5023
22.6712 | -35.7274 | -34.5087
22.6773 | -35.728 | -34.515
22.6834 | -35.7285 | -34.5215
22.6895 | -35.729 | -34.5274
22.6956 | -35.7296 | -34.5339
22.7017 | -35.7302 | -34.5401
22.7079 | -35.7307 | -34.5462
22.714 | -35.7313 | -34.5526
22.7201 | -35.7318 | -34.5589
22.7262 | -35.7324 | -34.5649
22.7323 | -35.733 | -34.5712
22.7384 | -35.7335 | -34.5775
22.7445 | -35.7341 | -34.5842
22.7506 | -35.7346 | -34.5904
22.7567 | -35.7352 | -34.5967
22.7628 | -35.7357 | -34.6026
22.7689 | -35.7362 | -34.6089
22.775 | -35.7368 | -34.6154
22.7811 | -35.7373 | -34.6216
22.7872 | -35.7379 | -34.628
22.7933 | -35.7384 | -34.6339
22.7994 | -35.739 | -34.6402
22.8055 | -35.7395 | -34.6467
22.8116 | -35.7401 | -34.653
22.8177 | -35.7406 | -34.6593
22.8238 | -35.7411 | -34.6657
22.8299 | -35.7417 | -34.6719
22.836 | -35.7423 | -34.6782
22.8421 | -35.7428 | -34.6846
22.8482 | -35.7433 | -34.691
22.8544 | -35.7439 | -34.6971
22.8605 | -35.7445 | -34.7035
22.8666 | -35.745 | -34.7093
22.8727 | -35.7455 | -34.7154
22.8788 | -35.7461 | -34.7216
22.8849 | -35.7466 | -34.7278
22.891 | -35.7472 | -34.734
22.8971 | -35.7477 | -34.7402
22.9032 | -35.7483 | -34.7463
22.9093 | -35.7488 | -34.7526
22.9154 | -35.7494 | -34.7588
22.9215 | -35.7499 | -34.7651
22.9276 | -35.7505 | -34.7713
22.9337 | -35.751 | -34.7774
22.9398 | -35.7516 | -34.7837
22.9459 | -35.7522 | -34.7902
22.952 | -35.7527 | -34.7965
22.9581 | -35.7532 | -34.8027
22.9642 | -35.7538 | -34.809
22.9703 | -35.7543 | -34.8149
22.9764 | -35.7549 | -34.8212
22.9825 | -35.7554 | -34.8276
22.9886 | -35.7559 | -34.8339
22.9948 | -35.7565 | -34.8402
23.0009 | -35.757 | -34.8466
23.007 | -35.7576 | -34.8528
23.0131 | -35.7581 | -34.8591
23.0192 | -35.7587 | -34.8654
23.0253 | -35.7592 | -34.8717
23.0314 | -35.7597 | -34.8782
23.0375 | -35.7603 | -34.8845
23.0436 | -35.7608 | -34.8906
23.0497 | -35.7614 | -34.8968
23.0558 | -35.7619 | -34.9029
23.0619 | -35.7624 | -34.9091
23.068 | -35.763 | -34.9154
23.0741 | -35.7635 | -34.9213
23.0802 | -35.764 | -34.9275
23.0863 | -35.7646 | -34.9337
23.0924 | -35.7651 | -34.9401
23.0985 | -35.7657 | -34.9462
23.1046 | -35.7662 | -34.9524
23.1107 | -35.7668 | -34.9586
23.1168 | -35.7673 | -34.965
23.1229 | -35.7679 | -34.9713
23.129 | -35.7684 | -34.9777
23.1351 | -35.769 | -34.9837
23.1413 | -35.7695 | -34.9901
23.1474 | -35.77 | -34.9961
23.1535 | -35.7706 | -35.0023
23.1596 | -35.7711 | -35.0086
23.1657 | -35.7717 | -35.0149
23.1718 | -35.7723 | -35.0208
23.1779 | -35.7728 | -35.0271
23.184 | -35.7733 | -35.0332
23.1901 | -35.7738 | -35.0393
23.1962 | -35.7744 | -35.0456
23.2023 | -35.775 | -35.0516
23.2084 | -35.7755 | -35.0579
23.2145 | -35.776 | -35.0638
23.2206 | -35.7766 | -35.07
23.2267 | -35.7771 | -35.0761
23.2328 | -35.7777 | -35.0823
23.2389 | -35.7782 | -35.0885
23.245 | -35.7787 | -35.0948
23.2511 | -35.7793 | -35.1007
23.2572 | -35.7798 | -35.1069
23.2633 | -35.7803 | -35.1133
23.2694 | -35.7809 | -35.1195
23.2755 | -35.7814 | -35.1258
23.2817 | -35.782 | -35.132
23.2878 | -35.7825 | -35.1381
23.2939 | -35.7831 | -35.1442
23.3 | -35.7836 | -35.1503
23.3061 | -35.7841 | -35.1564
23.3122 | -35.7847 | -35.1627
23.3183 | -35.7852 | -35.1686
23.3244 | -35.7858 | -35.1746
23.3305 | -35.7863 | -35.1805
23.3366 | -35.7869 | -35.1866
23.3427 | -35.7874 | -35.1925
23.3488 | -35.788 | -35.1987
23.3549 | -35.7885 | -35.2047
23.361 | -35.7891 | -35.2107
23.3671 | -35.7896 | -35.217
23.3732 | -35.7901 | -35.2231
23.3793 | -35.7907 | -35.2291
23.3854 | -35.7912 | -35.2354
23.3915 | -35.7918 | -35.2414
23.3976 | -35.7923 | -35.2476
23.4037 | -35.7929 | -35.2537
23.4098 | -35.7935 | -35.2598
23.4159 | -35.794 | -35.2661
23.422 | -35.7946 | -35.2723
23.4282 | -35.7951 | -35.2782
23.4343 | -35.7956 | -35.2844
23.4404 | -35.7961 | -35.2904
23.4465 | -35.7967 | -35.2966
23.4526 | -35.7972 | -35.303
23.4587 | -35.7977 | -35.3091
23.4648 | -35.7983 | -35.3149
23.4709 | -35.7988 | -35.3211
23.477 | -35.7994 | -35.3273
23.4831 | -35.7999 | -35.3333
23.4892 | -35.8004 | -35.3394
23.4953 | -35.801 | -35.3457
23.5014 | -35.8015 | -35.3517
23.5075 | -35.802 | -35.3579
23.5136 | -35.8026 | -35.3641
23.5197 | -35.8031 | -35.3703
23.5258 | -35.8036 | -35.3764
23.5319 | -35.8041 | -35.3821
23.538 | -35.8047 | -35.3884
23.5441 | -35.8052 | -35.3944
23.5502 | -35.8057 | -35.4006
23.5563 | -35.8063 | -35.4069
23.5624 | -35.8068 | -35.4131
23.5686 | -35.8073 | -35.4188
23.5747 | -35.8078 | -35.4248
23.5808 | -35.8084 | -35.4308
23.5869 | -35.8089 | -35.4368
23.593 | -35.8094 | -35.4429
23.5991 | -35.81 | -35.4493
23.6052 | -35.8105 | -35.455
23.6113 | -35.811 | -35.4611
23.6174 | -35.8115 | -35.4672
23.6235 | -35.8121 | -35.4734
23.6296 | -35.8126 | -35.4795
23.6357 | -35.8132 | -35.4856
23.6418 | -35.8137 | -35.4916
23.6479 | -35.8143 | -35.4975
23.654 | -35.8148 | -35.5035
23.6601 | -35.8153 | -35.5095
23.6662 | -35.8159 | -35.5157
23.6723 | -35.8164 | -35.5215
23.6784 | -35.8169 | -35.5277
23.6845 | -35.8174 | -35.534
23.6906 | -35.8179 | -35.5401
23.6967 | -35.8185 | -35.546
23.7028 | -35.819 | -35.5523
23.7089 | -35.8195 | -35.558
23.7151 | -35.8201 | -35.5642
23.7212 | -35.8206 | -35.5707
23.7273 | -35.8211 | -35.5768
23.7334 | -35.8217 | -35.5829
23.7395 | -35.8222 | -35.589
23.7456 | -35.8227 | -35.595
23.7517 | -35.8232 | -35.6015
23.7578 | -35.8238 | -35.6073
23.7639 | -35.8243 | -35.6133
23.77 | -35.8249 | -35.6195
23.7761 | -35.8254 | -35.6257
23.7822 | -35.8259 | -35.6314
23.7883 | -35.8264 | -35.6374
23.7944 | -35.827 | -35.6436
23.8005 | -35.8275 | -35.6496
23.8066 | -35.828 | -35.6558
23.8127 | -35.8285 | -35.6617
23.8188 | -35.8291 | -35.6677
23.8249 | -35.8297 | -35.6739
23.831 | -35.8302 | -35.68
23.8371 | -35.8307 | -35.6861
23.8432 | -35.8313 | -35.6922
23.8493 | -35.8318 | -35.6981
23.8555 | -35.8323 | -35.7043
23.8616 | -35.8329 | -35.7105
23.8677 | -35.8334 | -35.7165
23.8738 | -35.8339 | -35.7225
23.8799 | -35.8345 | -35.7287
23.886 | -35.835 | -35.7344
23.8921 | -35.8355 | -35.7405
23.8982 | -35.836 | -35.7465
23.9043 | -35.8365 | -35.7525
23.9104 | -35.8371 | -35.7584
23.9165 | -35.8376 | -35.7644
23.9226 | -35.8381 | -35.7702
23.9287 | -35.8387 | -35.7764
23.9348 | -35.8392 | -35.7825
23.9409 | -35.8397 | -35.7887
23.947 | -35.8403 | -35.7945
23.9531 | -35.8408 | -35.8005
23.9592 | -35.8413 | -35.8066
23.9653 | -35.8418 | -35.8126
23.9714 | -35.8423 | -35.8185
23.9775 | -35.8428 | -35.8246
23.9836 | -35.8434 | -35.8305
23.9897 | -35.8439 | -35.8365
23.9958 | -35.8444 | -35.8427
24.002 | -35.8449 | -35.8486
24.0081 | -35.8454 | -35.8547
24.0142 | -35.8459 | -35.8607
24.0203 | -35.8464 | -35.8669
24.0264 | -35.847 | -35.8729
24.0325 | -35.8475 | -35.8789
24.0386 | -35.848 | -35.885
24.0447 | -35.8485 | -35.8909
24.0508 | -35.8491 | -35.8969
24.0569 | -35.8496 | -35.9029
24.063 | -35.8501 | -35.9086
24.0691 | -35.8506 | -35.9146
24.0752 | -35.8512 | -35.9209
24.0813 | -35.8517 | -35.927
24.0874 | -35.8522 | -35.9332
24.0935 | -35.8527 | -35.9392
24.0996 | -35.8532 | -35.9452
24.1057 | -35.8537 | -35.9515
24.1118 | -35.8542 | -35.9577
24.1179 | -35.8548 | -35.964
24.124 | -35.8553 | -35.9699
24.1301 | -35.8558 | -35.9757
24.1362 | -35.8563 | -35.9818
24.1424 | -35.8569 | -35.9879
24.1485 | -35.8574 | -35.9941
24.1546 | -35.8579 | -35.9999
24.1607 | -35.8585 | -36.006
24.1668 | -35.859 | -36.0116
24.1729 | -35.8595 | -36.0175
24.179 | -35.86 | -36.0234
24.1851 | -35.8605 | -36.0296
24.1912 | -35.861 | -36.0356
24.1973 | -35.8616 | -36.0416
24.2034 | -35.8621 | -36.0475
24.2095 | -35.8626 | -36.0534
24.2156 | -35.8631 | -36.0596
24.2217 | -35.8636 | -36.0656
24.2278 | -35.8642 | -36.0715
24.2339 | -35.8647 | -36.0775
24.24 | -35.8652 | -36.0834
24.2461 | -35.8657 | -36.0893
24.2522 | -35.8662 | -36.0954
24.2583 | -35.8668 | -36.1014
24.2644 | -35.8673 | -36.1076
24.2705 | -35.8678 | -36.1135
24.2766 | -35.8683 | -36.1195
24.2827 | -35.8688 | -36.1254
24.2889 | -35.8693 | -36.1314
24.295 | -35.8698 | -36.1373
24.3011 | -35.8703 | -36.1432
24.3072 | -35.8709 | -36.149
24.3133 | -35.8714 | -36.155
24.3194 | -35.8719 | -36.1611
24.3255 | -35.8725 | -36.1673
24.3316 | -35.873 | -36.1733
24.3377 | -35.8735 | -36.1791
24.3438 | -35.874 | -36.1849
24.3499 | -35.8745 | -36.1909
24.356 | -35.875 | -36.1969
24.3621 | -35.8755 | -36.2031
24.3682 | -35.876 | -36.2091
24.3743 | -35.8765 | -36.215
24.3804 | -35.877 | -36.2206
24.3865 | -35.8774 | -36.2268
24.3926 | -35.8779 | -36.2324
24.3987 | -35.8785 | -36.2383
24.4048 | -35.879 | -36.2442
24.4109 | -35.8795 | -36.2502
24.417 | -35.88 | -36.2561
24.4231 | -35.8805 | -36.2619
24.4293 | -35.881 | -36.2681
24.4354 | -35.8815 | -36.2741
24.4415 | -35.882 | -36.2803
24.4476 | -35.8825 | -36.2858
24.4537 | -35.883 | -36.2917
24.4598 | -35.8835 | -36.2976
24.4659 | -35.884 | -36.3038
24.472 | -35.8845 | -36.3098
24.4781 | -35.8851 | -36.3159
24.4842 | -35.8855 | -36.3214
24.4903 | -35.8861 | -36.3275
24.4964 | -35.8866 | -36.3335
24.5025 | -35.8871 | -36.3394
24.5086 | -35.8876 | -36.3455
24.5147 | -35.8881 | -36.3514
24.5208 | -35.8886 | -36.3572
24.5269 | -35.8891 | -36.3632
24.533 | -35.8897 | -36.3692
24.5391 | -35.8902 | -36.375
24.5452 | -35.8907 | -36.3808
24.5513 | -35.8912 | -36.3866
24.5574 | -35.8918 | -36.3926
24.5635 | -35.8923 | -36.3985
24.5696 | -35.8928 | -36.4044
24.5758 | -35.8933 | -36.4104
24.5819 | -35.8938 | -36.4162
24.588 | -35.8943 | -36.4219
24.5941 | -35.8948 | -36.4278
24.6002 | -35.8953 | -36.4336
24.6063 | -35.8959 | -36.4396
24.6124 | -35.8964 | -36.4455
24.6185 | -35.8969 | -36.4515
24.6246 | -35.8974 | -36.4571
24.6307 | -35.8979 | -36.4632
24.6368 | -35.8985 | -36.4694
24.6429 | -35.899 | -36.4754
24.649 | -35.8995 | -36.4814
24.6551 | -35.9001 | -36.4872
24.6612 | -35.9006 | -36.4929
24.6673 | -35.9011 | -36.4989
24.6734 | -35.9016 | -36.5048
24.6795 | -35.9021 | -36.5106
24.6856 | -35.9026 | -36.5167
24.6917 | -35.9032 | -36.5223
24.6978 | -35.9037 | -36.5281
24.7039 | -35.9042 | -36.5341
24.71 | -35.9047 | -36.54
24.7162 | -35.9052 | -36.5461
24.7223 | -35.9057 | -36.552
24.7284 | -35.9062 | -36.5578
24.7345 | -35.9067 | -36.5638
24.7406 | -35.9072 | -36.5698
24.7467 | -35.9077 | -36.5757
24.7528 | -35.9083 | -36.5817
24.7589 | -35.9088 | -36.5874
24.765 | -35.9093 | -36.5928
24.7711 | -35.9099 | -36.5988
24.7772 | -35.9104 | -36.6048
24.7833 | -35.9109 | -36.6107
24.7894 | -35.9114 | -36.617
24.7955 | -35.9119 | -36.623
24.8016 | -35.9124 | -36.6286
24.8077 | -35.9129 | -36.6346
24.8138 | -35.9135 | -36.6409
24.8199 | -35.914 | -36.6466
24.826 | -35.9145 | -36.6526
24.8321 | -35.915 | -36.6584
24.8382 | -35.9155 | -36.6642
24.8443 | -35.916 | -36.67
24.8504 | -35.9165 | -36.676
24.8565 | -35.9171 | -36.682
24.8627 | -35.9176 | -36.6878
24.8688 | -35.9181 | -36.6933
24.8749 | -35.9186 | -36.6992
24.881 | -35.9191 | -36.7051
24.8871 | -35.9196 | -36.711
24.8932 | -35.9202 | -36.7168
24.8993 | -35.9206 | -36.7228
24.9054 | -35.9211 | -36.7285
24.9115 | -35.9217 | -36.7343
24.9176 | -35.9222 | -36.7401
24.9237 | -35.9227 | -36.746
24.9298 | -35.9231 | -36.7518
24.9359 | -35.9236 | -36.7578
24.942 | -35.9241 | -36.7634
24.9481 | -35.9246 | -36.7692
24.9542 | -35.9251 | -36.7751
24.9603 | -35.9256 | -36.781
24.9664 | -35.9262 | -36.7869
24.9725 | -35.9267 | -36.7927
24.9786 | -35.9272 | -36.7982
24.9847 | -35.9277 | -36.8042
24.9908 | -35.9282 | -36.8101
24.9969 | -35.9287 | -36.816
25.0031 | -35.9292 | -36.8218
25.0092 | -35.9297 | -36.8275
25.0153 | -35.9302 | -36.8334
25.0214 | -35.9307 | -36.8392
25.0275 | -35.9312 | -36.8453
25.0336 | -35.9317 | -36.8509
25.0397 | -35.9322 | -36.857
25.0458 | -35.9328 | -36.8632
25.0519 | -35.9333 | -36.8691
25.058 | -35.9338 | -36.8749
25.0641 | -35.9343 | -36.8808
25.0702 | -35.9348 | -36.8864
25.0763 | -35.9354 | -36.8921
25.0824 | -35.9358 | -36.8978
25.0885 | -35.9363 | -36.9036
25.0946 | -35.9369 | -36.9096
25.1007 | -35.9373 | -36.9155
25.1068 | -35.9379 | -36.9212
25.1129 | -35.9383 | -36.927
25.119 | -35.9389 | -36.9328
25.1251 | -35.9394 | -36.9388
25.1312 | -35.9399 | -36.9446
25.1373 | -35.9404 | -36.9501
25.1435 | -35.9409 | -36.9559
25.1496 | -35.9414 | -36.9613
25.1557 | -35.9419 | -36.9671
25.1618 | -35.9424 | -36.9728
25.1679 | -35.9429 | -36.9788
25.174 | -35.9434 | -36.9847
25.1801 | -35.9439 | -36.9905
25.1862 | -35.9444 | -36.9964
25.1923 | -35.9449 | -37.0021
25.1984 | -35.9454 | -37.008
25.2045 | -35.9459 | -37.0137
25.2106 | -35.9464 | -37.0198
25.2167 | -35.9469 | -37.0259
25.2228 | -35.9473 | -37.0316
25.2289 | -35.9478 | -37.0374
25.235 | -35.9483 | -37.0432
25.2411 | -35.9488 | -37.0492
25.2472 | -35.9493 | -37.0551
25.2533 | -35.9498 | -37.0611
25.2594 | -35.9503 | -37.0669
25.2655 | -35.9508 | -37.0728
25.2716 | -35.9513 | -37.0785
25.2777 | -35.9518 | -37.0842
25.2838 | -35.9523 | -37.0903
25.29 | -35.9528 | -37.0958
25.2961 | -35.9533 | -37.1015
25.3022 | -35.9538 | -37.1074
25.3083 | -35.9544 | -37.1134
25.3144 | -35.9548 | -37.1192
25.3205 | -35.9553 | -37.1252
25.3266 | -35.9558 | -37.1311
25.3327 | -35.9563 | -37.137
25.3388 | -35.9568 | -37.143
25.3449 | -35.9573 | -37.1487
25.351 | -35.9578 | -37.1543
25.3571 | -35.9583 | -37.1603
25.3632 | -35.9588 | -37.1657
25.3693 | -35.9593 | -37.1715
25.3754 | -35.9598 | -37.1773
25.3815 | -35.9603 | -37.1832
25.3876 | -35.9607 | -37.1889
25.3937 | -35.9613 | -37.1948
25.3998 | -35.9617 | -37.2003
25.4059 | -35.9622 | -37.2062
25.412 | -35.9627 | -37.2118
25.4181 | -35.9633 | -37.2177
25.4242 | -35.9637 | -37.2234
25.4304 | -35.9642 | -37.2292
25.4365 | -35.9647 | -37.2349
25.4426 | -35.9652 | -37.2408
25.4487 | -35.9657 | -37.2469
25.4548 | -35.9662 | -37.2528
25.4609 | -35.9667 | -37.2588
25.467 | -35.9672 | -37.2644
25.4731 | -35.9676 | -37.2701
25.4792 | -35.9681 | -37.2756
25.4853 | -35.9687 | -37.2812
25.4914 | -35.9691 | -37.2871
25.4975 | -35.9696 | -37.2927
25.5036 | -35.9702 | -37.2986
25.5097 | -35.9707 | -37.3044
25.5158 | -35.9712 | -37.3103
25.5219 | -35.9717 | -37.3161
25.528 | -35.9722 | -37.3221
25.5341 | -35.9727 | -37.3276
25.5402 | -35.9732 | -37.3335
25.5463 | -35.9737 | -37.3395
25.5524 | -35.9742 | -37.3453
25.5585 | -35.9747 | -37.3511
25.5646 | -35.9752 | -37.3569
25.5707 | -35.9757 | -37.3628
25.5769 | -35.9762 | -37.3685
25.583 | -35.9767 | -37.3743
25.5891 | -35.9772 | -37.3803
25.5952 | -35.9777 | -37.3863
25.6013 | -35.9782 | -37.3922
25.6074 | -35.9788 | -37.398
25.6135 | -35.9793 | -37.4039
25.6196 | -35.9797 | -37.41
25.6257 | -35.9802 | -37.4154
25.6318 | -35.9807 | -37.421
25.6379 | -35.9812 | -37.4268
25.644 | -35.9817 | -37.4323
25.6501 | -35.9822 | -37.4382
25.6562 | -35.9827 | -37.444
25.6623 | -35.9831 | -37.4496
25.6684 | -35.9837 | -37.4554
25.6745 | -35.9841 | -37.4612
25.6806 | -35.9846 | -37.4669
25.6867 | -35.9851 | -37.4728
25.6928 | -35.9856 | -37.4787
25.6989 | -35.986 | -37.4846
25.705 | -35.9865 | -37.4905
25.7111 | -35.987 | -37.4962
25.7173 | -35.9875 | -37.5016
25.7234 | -35.988 | -37.5074
25.7295 | -35.9885 | -37.5132
25.7356 | -35.989 | -37.5192
25.7417 | -35.9895 | -37.5249
25.7478 | -35.99 | -37.5309
25.7539 | -35.9905 | -37.5365
25.76 | -35.991 | -37.5423
25.7661 | -35.9915 | -37.5482
25.7722 | -35.9919 | -37.5539
25.7783 | -35.9924 | -37.5596
25.7844 | -35.993 | -37.5654
25.7905 | -35.9935 | -37.5711
25.7966 | -35.9939 | -37.5768
25.8027 | -35.9943 | -37.5826
25.8088 | -35.9948 | -37.5885
25.8149 | -35.9953 | -37.5939
25.821 | -35.9957 | -37.5993
25.8271 | -35.9962 | -37.6049
25.8332 | -35.9967 | -37.6104
25.8393 | -35.9972 | -37.6163
25.8454 | -35.9977 | -37.6223
25.8515 | -35.9982 | -37.628
25.8576 | -35.9986 | -37.6334
25.8638 | -35.9991 | -37.6391
25.8699 | -35.9996 | -37.6448
25.876 | -36.0001 | -37.6503
25.8821 | -36.0006 | -37.6559
25.8882 | -36.0011 | -37.6616
25.8943 | -36.0016 | -37.6675
25.9004 | -36.0021 | -37.6733
25.9065 | -36.0026 | -37.6792
25.9126 | -36.0031 | -37.685
25.9187 | -36.0036 | -37.6904
25.9248 | -36.004 | -37.6959
25.9309 | -36.0045 | -37.7015
25.937 | -36.005 | -37.7072
25.9431 | -36.0055 | -37.7129
25.9492 | -36.006 | -37.7186
25.9553 | -36.0065 | -37.7246
25.9614 | -36.0069 | -37.73
25.9675 | -36.0074 | -37.7357
25.9736 | -36.0079 | -37.7413
25.9797 | -36.0084 | -37.7469
25.9858 | -36.0089 | -37.7527
25.9919 | -36.0094 | -37.7585
25.998 | -36.0098 | -37.7638
26.0042 | -36.0103 | -37.7697
26.0103 | -36.0108 | -37.7751
26.0164 | -36.0113 | -37.7809
26.0225 | -36.0118 | -37.7865
26.0286 | -36.0123 | -37.7923
26.0347 | -36.0128 | -37.7978
26.0408 | -36.0132 | -37.8035
26.0469 | -36.0137 | -37.8091
26.053 | -36.0142 | -37.8148
26.0591 | -36.0147 | -37.8206
26.0652 | -36.0152 | -37.8259
26.0713 | -36.0157 | -37.8315
26.0774 | -36.0162 | -37.8372
26.0835 | -36.0167 | -37.843
26.0896 | -36.0171 | -37.8487
26.0957 | -36.0176 | -37.8544
26.1018 | -36.0181 | -37.8601
26.1079 | -36.0185 | -37.8657
26.114 | -36.019 | -37.8713
26.1201 | -36.0194 | -37.8769
26.1262 | -36.0199 | -37.8824
26.1323 | -36.0204 | -37.8883
26.1384 | -36.0209 | -37.8936
26.1445 | -36.0214 | -37.8995
26.1507 | -36.0219 | -37.9053
26.1568 | -36.0224 | -37.911
26.1629 | -36.0228 | -37.9167
26.169 | -36.0234 | -37.9225
26.1751 | -36.0238 | -37.9279
26.1812 | -36.0243 | -37.9334
26.1873 | -36.0248 | -37.939
26.1934 | -36.0253 | -37.9447
26.1995 | -36.0257 | -37.9504
26.2056 | -36.0262 | -37.9562
26.2117 | -36.0267 | -37.9619
26.2178 | -36.0272 | -37.9674
26.2239 | -36.0276 | -37.9731
26.23 | -36.0281 | -37.9788
26.2361 | -36.0286 | -37.9843
26.2422 | -36.029 | -37.9895
26.2483 | -36.0295 | -37.9952
26.2544 | -36.03 | -38.0006
26.2605 | -36.0305 | -38.0063
26.2666 | -36.031 | -38.0121
26.2727 | -36.0314 | -38.0178
26.2788 | -36.0319 | -38.0231
26.2849 | -36.0324 | -38.0289
26.2911 | -36.0329 | -38.0345
26.2972 | -36.0334 | -38.0403
26.3033 | -36.0339 | -38.0458
26.3094 | -36.0344 | -38.0516
26.3155 | -36.0349 | -38.057
26.3216 | -36.0354 | -38.0629
26.3277 | -36.0358 | -38.0684
26.3338 | -36.0363 | -38.0741
26.3399 | -36.0368 | -38.08
26.346 | -36.0373 | -38.0853
26.3521 | -36.0378 | -38.0913
26.3582 | -36.0382 | -38.0969
26.3643 | -36.0387 | -38.1023
26.3704 | -36.0392 | -38.1077
26.3765 | -36.0397 | -38.1132
26.3826 | -36.0402 | -38.1186
26.3887 | -36.0406 | -38.1245
26.3948 | -36.0411 | -38.1299
26.4009 | -36.0416 | -38.1357
26.407 | -36.042 | -38.1413
26.4131 | -36.0425 | -38.1469
26.4192 | -36.043 | -38.1523
26.4253 | -36.0434 | -38.158
26.4314 | -36.0439 | -38.1635
26.4376 | -36.0444 | -38.1692
26.4437 | -36.0449 | -38.1751
26.4498 | -36.0454 | -38.1808
26.4559 | -36.0458 | -38.186
26.462 | -36.0463 | -38.1914
26.4681 | -36.0467 | -38.1973
26.4742 | -36.0472 | -38.2027
26.4803 | -36.0476 | -38.2086
26.4864 | -36.0481 | -38.2136
26.4925 | -36.0486 | -38.2192
26.4986 | -36.0491 | -38.2248
26.5047 | -36.0496 | -38.2303
26.5108 | -36.0501 | -38.2358
26.5169 | -36.0505 | -38.2416
26.523 | -36.051 | -38.2475
26.5291 | -36.0515 | -38.253
26.5352 | -36.0519 | -38.2588
26.5413 | -36.0523 | -38.2644
26.5474 | -36.0528 | -38.27
26.5535 | -36.0533 | -38.2757
26.5596 | -36.0538 | -38.281
26.5657 | -36.0542 | -38.2866
26.5718 | -36.0547 | -38.2923
26.578 | -36.0552 | -38.298
26.5841 | -36.0557 | -38.3033
26.5902 | -36.0562 | -38.3089
26.5963 | -36.0567 | -38.3143
26.6024 | -36.0571 | -38.3201
26.6085 | -36.0575 | -38.326
26.6146 | -36.058 | -38.3318
26.6207 | -36.0585 | -38.3374
26.6268 | -36.059 | -38.3432
26.6329 | -36.0594 | -38.3487
26.639 | -36.0599 | -38.3543
26.6451 | -36.0604 | -38.3599
26.6512 | -36.0608 | -38.3654
26.6573 | -36.0613 | -38.3708
26.6634 | -36.0618 | -38.3759
26.6695 | -36.0622 | -38.3816
26.6756 | -36.0627 | -38.3875
26.6817 | -36.0632 | -38.3932
26.6878 | -36.0637 | -38.3986
26.6939 | -36.0641 | -38.4043
26.7 | -36.0645 | -38.4094
26.7061 | -36.065 | -38.4149
26.7122 | -36.0655 | -38.4204
26.7183 | -36.0659 | -38.426
26.7245 | -36.0664 | -38.4314
26.7306 | -36.0668 | -38.4371
26.7367 | -36.0672 | -38.4424
26.7428 | -36.0677 | -38.4479
26.7489 | -36.0682 | -38.4536
26.755 | -36.0686 | -38.4586
26.7611 | -36.0691 | -38.4643
26.7672 | -36.0695 | -38.4698
26.7733 | -36.07 | -38.4753
26.7794 | -36.0704 | -38.4808
26.7855 | -36.0709 | -38.4865
26.7916 | -36.0713 | -38.4919
26.7977 | -36.0718 | -38.4972
26.8038 | -36.0723 | -38.5026
26.8099 | -36.0728 | -38.5081
26.816 | -36.0733 | -38.5135
26.8221 | -36.0737 | -38.5189
26.8282 | -36.0742 | -38.5244
26.8343 | -36.0747 | -38.5299
26.8404 | -36.0751 | -38.535
26.8465 | -36.0756 | -38.5406
26.8526 | -36.0761 | -38.5463
26.8587 | -36.0765 | -38.5517
26.8649 | -36.0769 | -38.5573
26.871 | -36.0774 | -38.5628
26.8771 | -36.0779 | -38.5682
26.8832 | -36.0784 | -38.5737
26.8893 | -36.0788 | -38.5792
26.8954 | -36.0793 | -38.5844
26.9015 | -36.0797 | -38.59
26.9076 | -36.0802 | -38.5956
26.9137 | -36.0807 | -38.6008
26.9198 | -36.0812 | -38.6066
26.9259 | -36.0816 | -38.6122
26.932 | -36.082 | -38.6172
26.9381 | -36.0824 | -38.6227
26.9442 | -36.0829 | -38.6278
26.9503 | -36.0833 | -38.6334
26.9564 | -36.0838 | -38.6388
26.9625 | -36.0843 | -38.6444
26.9686 | -36.0848 | -38.65
26.9747 | -36.0852 | -38.6553
26.9808 | -36.0857 | -38.6608
26.9869 | -36.0862 | -38.6665
26.993 | -36.0866 | -38.6721
26.9991 | -36.0871 | -38.6777
27.0052 | -36.0875 | -38.6833
27.0114 | -36.088 | -38.6891
27.0175 | -36.0885 | -38.6947
27.0236 | -36.0889 | -38.7006
27.0297 | -36.0894 | -38.7062
27.0358 | -36.0899 | -38.7115
27.0419 | -36.0904 | -38.7171
27.048 | -36.0908 | -38.7226
27.0541 | -36.0913 | -38.728
27.0602 | -36.0917 | -38.7333
27.0663 | -36.0922 | -38.7389
27.0724 | -36.0927 | -38.744
27.0785 | -36.0931 | -38.7494
27.0846 | -36.0936 | -38.7544
27.0907 | -36.094 | -38.7601
27.0968 | -36.0945 | -38.7658
27.1029 | -36.0949 | -38.7711
27.109 | -36.0954 | -38.7763
27.1151 | -36.0958 | -38.7818
27.1212 | -36.0963 | -38.7866
27.1273 | -36.0967 | -38.7922
27.1334 | -36.0972 | -38.7977
27.1395 | -36.0977 | -38.8034
27.1456 | -36.0982 | -38.8086
27.1518 | -36.0987 | -38.8139
27.1579 | -36.0991 | -38.8197
27.164 | -36.0996 | -38.8252
27.1701 | -36.1001 | -38.8308
27.1762 | -36.1006 | -38.8365
27.1823 | -36.101 | -38.8424
27.1884 | -36.1015 | -38.8474
27.1945 | -36.1019 | -38.8524
27.2006 | -36.1023 | -38.8577
27.2067 | -36.1028 | -38.8628
27.2128 | -36.1032 | -38.8684
27.2189 | -36.1037 | -38.8737
27.225 | -36.1042 | -38.8789
27.2311 | -36.1046 | -38.884
27.2372 | -36.1051 | -38.8896
27.2433 | -36.1055 | -38.8948
27.2494 | -36.106 | -38.9004
27.2555 | -36.1065 | -38.906
27.2616 | -36.107 | -38.9113
27.2677 | -36.1074 | -38.9168
27.2738 | -36.1078 | -38.9222
27.2799 | -36.1083 | -38.9279
27.286 | -36.1087 | -38.9331
27.2921 | -36.1092 | -38.9388
27.2983 | -36.1096 | -38.9443
27.3044 | -36.1101 | -38.9499
27.3105 | -36.1105 | -38.9554
27.3166 | -36.111 | -38.9609
27.3227 | -36.1115 | -38.9664
27.3288 | -36.1119 | -38.9718
27.3349 | -36.1123 | -38.9775
27.341 | -36.1128 | -38.9831
27.3471 | -36.1132 | -38.9885
27.3532 | -36.1137 | -38.9936
27.3593 | -36.1141 | -38.9988
27.3654 | -36.1146 | -39.0042
27.3715 | -36.115 | -39.0088
27.3776 | -36.1154 | -39.0144
27.3837 | -36.1158 | -39.0198
27.3898 | -36.1163 | -39.025
27.3959 | -36.1167 | -39.0306
27.402 | -36.1172 | -39.0359
27.4081 | -36.1176 | -39.0411
27.4142 | -36.1181 | -39.0464
27.4203 | -36.1185 | -39.0517
27.4264 | -36.1189 | -39.057
27.4325 | -36.1194 | -39.0622
27.4387 | -36.1199 | -39.0671
27.4448 | -36.1203 | -39.0723
27.4509 | -36.1207 | -39.0778
27.457 | -36.1212 | -39.0834
27.4631 | -36.1216 | -39.0884
27.4692 | -36.1221 | -39.094
27.4753 | -36.1225 | -39.099
27.4814 | -36.123 | -39.1043
27.4875 | -36.1234 | -39.1094
27.4936 | -36.1239 | -39.1146
27.4997 | -36.1243 | -39.1202
27.5058 | -36.1247 | -39.1255
27.5119 | -36.1251 | -39.1303
27.518 | -36.1256 | -39.1359
27.5241 | -36.126 | -39.1414
27.5302 | -36.1265 | -39.147
27.5363 | -36.1271 | -39.1524
27.5424 | -36.1274 | -39.1578
27.5485 | -36.1279 | -39.1635
27.5546 | -36.1284 | -39.169
27.5607 | -36.1289 | -39.1742
27.5668 | -36.1293 | -39.1795
27.5729 | -36.1298 | -39.185
27.5791 | -36.1302 | -39.1897
27.5852 | -36.1307 | -39.1952
27.5913 | -36.1311 | -39.2003
27.5974 | -36.1316 | -39.2058
27.6035 | -36.132 | -39.2113
27.6096 | -36.1324 | -39.2165
27.6157 | -36.1329 | -39.2218
27.6218 | -36.1333 | -39.2272
27.6279 | -36.1338 | -39.2326
27.634 | -36.1342 | -39.238
27.6401 | -36.1346 | -39.2432
27.6462 | -36.1351 | -39.2484
27.6523 | -36.1355 | -39.2535
27.6584 | -36.136 | -39.2591
27.6645 | -36.1365 | -39.2646
27.6706 | -36.137 | -39.2701
27.6767 | -36.1374 | -39.2752
27.6828 | -36.1379 | -39.2807
27.6889 | -36.1383 | -39.2859
27.695 | -36.1388 | -39.291
27.7011 | -36.1393 | -39.2963
27.7072 | -36.1397 | -39.3015
27.7133 | -36.1402 | -39.3072
27.7194 | -36.1406 | -39.3121
27.7256 | -36.1411 | -39.3173
27.7317 | -36.1415 | -39.3223
27.7378 | -36.142 | -39.3281
27.7439 | -36.1424 | -39.3335
27.75 | -36.1429 | -39.339
27.7561 | -36.1432 | -39.3441
27.7622 | -36.1436 | -39.3497
27.7683 | -36.1441 | -39.355
27.7744 | -36.1446 | -39.3602
27.7805 | -36.145 | -39.3657
27.7866 | -36.1455 | -39.371
27.7927 | -36.1459 | -39.376
27.7988 | -36.1464 | -39.3811
27.8049 | -36.1469 | -39.3864
27.811 | -36.1473 | -39.3917
27.8171 | -36.1477 | -39.3968
27.8232 | -36.1482 | -39.4021
27.8293 | -36.1486 | -39.4075
27.8354 | -36.1491 | -39.413
27.8415 | -36.1496 | -39.418
27.8476 | -36.15 | -39.4235
27.8537 | -36.1504 | -39.4287
27.8598 | -36.1508 | -39.4339
27.866 | -36.1512 | -39.4389
27.8721 | -36.1517 | -39.4445
27.8782 | -36.1522 | -39.4496
27.8843 | -36.1526 | -39.4547
27.8904 | -36.153 | -39.4599
27.8965 | -36.1534 | -39.4645
27.9026 | -36.1539 | -39.4699
27.9087 | -36.1542 | -39.4752
27.9148 | -36.1546 | -39.4803
27.9209 | -36.1551 | -39.4855
27.927 | -36.1555 | -39.4905
27.9331 | -36.1559 | -39.4955
27.9392 | -36.1563 | -39.5012
27.9453 | -36.1568 | -39.5064
27.9514 | -36.1571 | -39.5118
27.9575 | -36.1576 | -39.5177
27.9636 | -36.158 | -39.5229
27.9697 | -36.1584 | -39.5285
27.9758 | -36.1588 | -39.5339
27.9819 | -36.1591 | -39.539
27.988 | -36.1596 | -39.5441
27.9941 | -36.16 | -39.5493
28.0002 | -36.1604 | -39.5543
28.0063 | -36.1609 | -39.5595
28.0125 | -36.1613 | -39.5648
28.0186 | -36.1618 | -39.5704
28.0247 | -36.1622 | -39.5758
28.0308 | -36.1626 | -39.5809
28.0369 | -36.163 | -39.5862
28.043 | -36.1635 | -39.5916
28.0491 | -36.1639 | -39.597
28.0552 | -36.1643 | -39.6024
28.0613 | -36.1648 | -39.6074
28.0674 | -36.1652 | -39.613
28.0735 | -36.1655 | -39.618

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.9493 | -44.0591
21.4908 | -35.8949 | -36.5987
16.4948 | -35.2 | -28.9969
12.6602 | -34.7914 | -22.5393
9.7171 | -34.5477 | -17.472
7.4582 | -34.375 | -13.3744
5.7244 | -34.2534 | -10.0491
4.3936 | -34.1715 | -7.3593
3.3722 | -34.1434 | -5.2469
2.5883 | -34.1182 | -3.6263
1.9866 | -34.1135 | -2.338
1.5248 | -34.1085 | -1.304
1.1703 | -34.1166 | -0.62794
1 | -34.0993 | -0.27462
0.89822 | -34.0988 | 0.0078344
0.68941 | -34.0605 | 0.60318
0.52919 | -34.0701 | 1.0812
0.40612 | -34.0784 | 1.6963
0.31171 | -34.0392 | 2.1813
0.23928 | -34.0604 | 2.916
0.18358 | -34.0672 | 3.3383
0.14091 | -34.0536 | 4.2411
0.10819 | -34.0815 | 5.3998
0.083017 | -34.0904 | 6.8895
0.063705 | -34.1561 | 8.7599
0.048923 | -34.2089 | 11.2542
0.037479 | -34.3113 | 14.4554
0.028801 | -34.4905 | 18.4759
0.022078 | -34.7879 | 23.8806
0.017023 | -35.2708 | 30.319
0.013018 | -36.0774 | 38.0023
0.010014 | -37.1814 | 45.9369

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003689 | -2.2333
0.007377 | -27.2604
0.011066 | -35.1904
0.014754 | -40.3363
0.018443 | -43.6262
0.022131 | -45.8652
0.02582 | -47.226
0.029508 | -47.6675
0.033197 | -49.7691
0.036885 | -52.0955
0.040574 | -53.1584
0.044262 | -52.7284
0.047951 | -53.0399
0.051639 | -53.7646
0.055328 | -55.2148
0.059016 | -57.1096
0.062705 | -57.7219
0.066393 | -57.491
0.070082 | -58.7754
0.07377 | -60.1054
0.077459 | -59.6254
0.081147 | -60.8942
0.084836 | -62.7148
0.088524 | -62.9076
0.092213 | -62.8125
0.095901 | -63.4171
0.09959 | -63.7017
0.10328 | -64.2006
0.10697 | -64.9656
0.11066 | -65.3961
0.11434 | -64.8125
0.11803 | -65.3454
0.12172 | -65.7071
0.12541 | -65.8605
0.1291 | -66.6041
0.13279 | -67.6616
0.13648 | -67.2738
0.14016 | -67.2545
0.14385 | -66.2031
0.14754 | -66.4576
0.15123 | -67.6166
0.15492 | -68.908
0.15861 | -69.443
0.16229 | -68.0651
0.16598 | -68.0974
0.16967 | -69.1152
0.17336 | -69.5232
0.17705 | -70.071
0.18074 | -69.8295
0.18443 | -70.0702
0.18811 | -69.8525
0.1918 | -70.5002
0.19549 | -70.4102
0.19918 | -70.4625
0.20287 | -70.7014
0.20656 | -70.207
0.21024 | -70.7997
0.21393 | -71.1039
0.21762 | -70.8362
0.22131 | -70.6925
0.225 | -71.5179
0.22869 | -72.1887
0.23238 | -71.8708
0.23606 | -72.9956
0.23975 | -72.5104
0.24344 | -72.5174
0.24713 | -73.2548
0.25082 | -73.063
0.25451 | -71.9813
0.2582 | -73.2821
0.26188 | -74.3984
0.26557 | -74.2282
0.26926 | -73.7755
0.27295 | -73.6981
0.27664 | -74.7317
0.28033 | -74.7369
0.28402 | -74.559
0.2877 | -74.832
0.29139 | -74.8566
0.29508 | -75.3781
0.29877 | -74.9366
0.30246 | -74.0204
0.30615 | -73.889
0.30983 | -73.7679
0.31352 | -74.7065
0.31721 | -75.5391
0.3209 | -75.37
0.32459 | -74.9327
0.32828 | -74.8547
0.33197 | -75.1439
0.33565 | -75.5624
0.33934 | -76.1238
0.34303 | -75.7757
0.34672 | -75.3695
0.35041 | -75.9611
0.3541 | -76.2553
0.35779 | -77.7737
0.36147 | -77.2925
0.36516 | -76.1337
0.36885 | -75.592
0.37254 | -75.4173
0.37623 | -75.9276
0.37992 | -76.8788
0.3836 | -77.5456
0.38729 | -77.5322
0.39098 | -77.656
0.39467 | -77.2844
0.39836 | -77.7026
0.40205 | -77.3985
0.40574 | -77.391
0.40942 | -76.811
0.41311 | -77.0258
0.4168 | -77.0095
0.42049 | -76.7715
0.42418 | -77.1606
0.42787 | -77.3856
0.43156 | -77.7969
0.43524 | -77.8868
0.43893 | -76.8918
0.44262 | -77.7517
0.44631 | -78.0971
0.45 | -78.3798
0.45369 | -78.4609
0.45737 | -78.6497
0.46106 | -79.1584
0.46475 | -79.0937
0.46844 | -79.2916
0.47213 | -78.2336
0.47582 | -78.0419
0.47951 | -77.9784
0.48319 | -77.7729
0.48688 | -78.3589
0.49057 | -79.3386
0.49426 | -78.7239
0.49795 | -78.6541
0.50164 | -79.4597
0.50533 | -79.3598
0.50901 | -79.5667
0.5127 | -78.9091
0.51639 | -80.2107
0.52008 | -80.3102
0.52377 | -80.3835
0.52746 | -79.3758
0.53114 | -79.1057
0.53483 | -79.521
0.53852 | -79.9601
0.54221 | -80.0065
0.5459 | -79.6631
0.54959 | -79.2923
0.55328 | -79.3695
0.55696 | -79.713
0.56065 | -79.8851
0.56434 | -79.7596
0.56803 | -79.93
0.57172 | -80.4839
0.57541 | -81.0981
0.5791 | -81.427
0.58278 | -81.4236
0.58647 | -81.2687
0.59016 | -81.3412
0.59385 | -81.4018
0.59754 | -81.0172
0.60123 | -80.5944
0.60491 | -80.7891
0.6086 | -81.1588
0.61229 | -81.0671
0.61598 | -80.8287
0.61967 | -81.0882
0.62336 | -81.9174
0.62705 | -82.3648
0.63073 | -82.1411
0.63442 | -82.1052
0.63811 | -82.2052
0.6418 | -82.4259
0.64549 | -82.6257
0.64918 | -82.325
0.65287 | -81.8531
0.65655 | -81.887
0.66024 | -82.0197
0.66393 | -82.0737
0.66762 | -81.916
0.67131 | -81.6289
0.675 | -81.5271
0.67868 | -81.5843
0.68237 | -81.8063
0.68606 | -81.8488
0.68975 | -81.7713
0.69344 | -82.007
0.69713 | -82.3816
0.70082 | -82.5908
0.70451 | -82.9713
0.70819 | -83.3145
0.71188 | -83.3898
0.71557 | -83.3649
0.71926 | -83.7279
0.72295 | -83.8595
0.72664 | -83.374
0.73032 | -83.16
0.73401 | -83.1893
0.7377 | -83.0333
0.74139 | -82.7521
0.74508 | -82.4953
0.74877 | -82.5747
0.75245 | -82.915
0.75614 | -82.7848
0.75983 | -82.6152
0.76352 | -82.9609
0.76721 | -83.451
0.7709 | -83.6066
0.77459 | -83.7999
0.77828 | -83.5108
0.78196 | -83.2155
0.78565 | -83.7422
0.78934 | -84.503
0.79303 | -84.7874
0.79672 | -84.7358
0.80041 | -84.6574
0.80409 | -84.2739
0.80778 | -83.7707
0.81147 | -83.8674
0.81516 | -84.2527
0.81885 | -84.2974
0.82254 | -84.1398
0.82622 | -84.1262
0.82991 | -84.1789
0.8336 | -84.3264
0.83729 | -84.57
0.84098 | -84.3809
0.84467 | -84.2429
0.84836 | -84.6579
0.85205 | -85.1215
0.85573 | -84.9798
0.85942 | -84.633
0.86311 | -84.7895
0.8668 | -85.1935
0.87049 | -85.5795
0.87418 | -85.4664
0.87786 | -84.8372
0.88155 | -84.5483
0.88524 | -85.2071
0.88893 | -86.0547
0.89262 | -86.0388
0.89631 | -85.6606
0.89999 | -85.284
0.90368 | -85.3144
0.90737 | -85.5977
0.91106 | -85.472
0.91475 | -84.9949
0.91844 | -84.8715
0.92213 | -85.0476
0.92582 | -85.0084
0.9295 | -84.8834
0.93319 | -84.7892
0.93688 | -84.9759
0.94057 | -85.3912
0.94426 | -85.8899
0.94795 | -86.0173
0.95163 | -85.7136
0.95532 | -85.8053
0.95901 | -85.7986
0.9627 | -85.5079
0.96639 | -85.1967
0.97008 | -85.0729
0.97377 | -85.5667
0.97745 | -86.1549
0.98114 | -86.363
0.98483 | -86.3342
0.98852 | -86.4226
0.99221 | -86.6853
0.9959 | -86.5839
0.99958 | -86.4157
1.0033 | -86.4323
1.007 | -86.358
1.0107 | -86.079
1.0143 | -85.9522
1.018 | -85.6393
1.0217 | -85.4195
1.0254 | -85.6652
1.0291 | -85.9665
1.0328 | -86.3506
1.0365 | -87.0922
1.0402 | -87.1517
1.0438 | -86.4852
1.0475 | -86.2605
1.0512 | -86.4553
1.0549 | -86.7164
1.0586 | -86.6789
1.0623 | -86.8256
1.066 | -86.8708
1.0697 | -86.7656
1.0734 | -86.5849
1.077 | -86.4816
1.0807 | -86.569
1.0844 | -86.7737
1.0881 | -86.9929
1.0918 | -87.2871
1.0955 | -87.6666
1.0992 | -87.862
1.1029 | -87.6113
1.1066 | -87.1555
1.1102 | -86.8289
1.1139 | -86.6501
1.1176 | -86.6261
1.1213 | -86.8527
1.125 | -87.2062
1.1287 | -87.3233
1.1324 | -87.0872
1.1361 | -86.8199
1.1397 | -86.8329
1.1434 | -87.1292
1.1471 | -87.5407
1.1508 | -87.9305
1.1545 | -88.2303
1.1582 | -88.3408
1.1619 | -88.1743
1.1656 | -87.9286
1.1693 | -87.8742
1.1729 | -88.0155
1.1766 | -88.1509
1.1803 | -88.2136
1.184 | -88.322
1.1877 | -88.4861
1.1914 | -88.6313
1.1951 | -88.7603
1.1988 | -88.9758
1.2025 | -89.2115
1.2061 | -89.1546
1.2098 | -88.8482
1.2135 | -88.6666
1.2172 | -88.6084
1.2209 | -88.4661
1.2246 | -88.3339
1.2283 | -88.384
1.232 | -88.5283
1.2356 | -88.5638
1.2393 | -88.5173
1.243 | -88.5943
1.2467 | -88.6305
1.2504 | -88.3454
1.2541 | -88.0437
1.2578 | -88.0223
1.2615 | -88.0924
1.2652 | -88.1489
1.2688 | -88.3333
1.2725 | -88.6059
1.2762 | -88.8827
1.2799 | -89.2494
1.2836 | -89.5756
1.2873 | -89.6048
1.291 | -89.3564
1.2947 | -88.9927
1.2984 | -88.6164
1.302 | -88.3797
1.3057 | -88.4417
1.3094 | -88.7156
1.3131 | -88.9122
1.3168 | -88.922
1.3205 | -88.8442
1.3242 | -88.9011
1.3279 | -89.1232
1.3316 | -89.2045
1.3352 | -89.0791
1.3389 | -89.0513
1.3426 | -89.2087
1.3463 | -89.348
1.35 | -89.3211
1.3537 | -89.1557
1.3574 | -88.9149
1.3611 | -88.7336
1.3647 | -88.7606
1.3684 | -88.9957
1.3721 | -89.3219
1.3758 | -89.6244
1.3795 | -89.824
1.3832 | -89.8702
1.3869 | -89.7839
1.3906 | -89.7289
1.3943 | -89.7913
1.3979 | -89.8135
1.4016 | -89.7267
1.4053 | -89.6579
1.409 | -89.6446
1.4127 | -89.6386
1.4164 | -89.6268
1.4201 | -89.5276
1.4238 | -89.391
1.4275 | -89.4596
1.4311 | -89.8027
1.4348 | -90.1383
1.4385 | -90.1082
1.4422 | -89.8564
1.4459 | -89.7422
1.4496 | -89.6736
1.4533 | -89.557
1.457 | -89.6492
1.4606 | -89.9949
1.4643 | -90.225
1.468 | -90.1344
1.4717 | -89.9472
1.4754 | -89.8656
1.4791 | -89.8616
1.4828 | -89.8235
1.4865 | -89.769
1.4902 | -89.7814
1.4938 | -89.8858
1.4975 | -89.9883
1.5012 | -89.957
1.5049 | -89.8089
1.5086 | -89.5981
1.5123 | -89.3235
1.516 | -89.1568
1.5197 | -89.2615
1.5234 | -89.494
1.527 | -89.5944
1.5307 | -89.6244
1.5344 | -89.7715
1.5381 | -90.0184
1.5418 | -90.1563
1.5455 | -90.1463
1.5492 | -90.2118
1.5529 | -90.4515
1.5565 | -90.6933
1.5602 | -90.7521
1.5639 | -90.7182
1.5676 | -90.7851
1.5713 | -90.9381
1.575 | -91.0451
1.5787 | -90.9649
1.5824 | -90.6148
1.5861 | -90.155
1.5897 | -89.8687
1.5934 | -89.9487
1.5971 | -90.0018
1.6008 | -90.0782
1.6045 | -90.1308
1.6082 | -90.1548
1.6119 | -90.1605
1.6156 | -90.1874
1.6193 | -90.2853
1.6229 | -90.4647
1.6266 | -90.703
1.6303 | -90.9619
1.634 | -91.1582
1.6377 | -91.2274
1.6414 | -91.2117
1.6451 | -91.1878
1.6488 | -91.1596
1.6525 | -91.0843
1.6561 | -90.9715
1.6598 | -90.8816
1.6635 | -90.8246
1.6672 | -90.7588
1.6709 | -90.6841
1.6746 | -90.6625
1.6783 | -90.7458
1.682 | -90.9051
1.6856 | -91.1103
1.6893 | -91.3658
1.693 | -91.6307
1.6967 | -91.8155
1.7004 | -91.8184
1.7041 | -91.6582
1.7078 | -91.4787
1.7115 | -91.4038
1.7152 | -91.4592
1.7188 | -91.5874
1.7225 | -91.7007
1.7262 | -91.7134
1.7299 | -91.5853
1.7336 | -91.4003
1.7373 | -91.318
1.741 | -91.3915
1.7447 | -91.5537
1.7484 | -91.709
1.752 | -91.7964
1.7557 | -91.8022
1.7594 | -91.756
1.7631 | -91.7228
1.7668 | -91.729
1.7705 | -91.7359
1.7742 | -91.7382
1.7779 | -91.7392
1.7815 | -91.7295
1.7852 | -91.7019
1.7889 | -91.7092
1.7926 | -91.8143
1.7963 | -91.971
1.8 | -92.022
1.8037 | -91.8671
1.8074 | -91.6199
1.8111 | -91.4158
1.8147 | -91.2933
1.8184 | -91.2406
1.8221 | -91.2731
1.8258 | -91.3726
1.8295 | -91.491
1.8332 | -91.6176
1.8369 | -91.7588
1.8406 | -91.8749
1.8443 | -91.8978
1.8479 | -91.7954
1.8516 | -91.6648
1.8553 | -91.5862
1.859 | -91.5403
1.8627 | -91.4576
1.8664 | -91.3593
1.8701 | -91.3175
1.8738 | -91.3279
1.8774 | -91.3331
1.8811 | -91.3108
1.8848 | -91.3186
1.8885 | -91.4581
1.8922 | -91.7452
1.8959 | -92.0521
1.8996 | -92.2384
1.9033 | -92.2751
1.907 | -92.1978
1.9106 | -92.0486
1.9143 | -91.9011
1.918 | -91.8434
1.9217 | -91.9311
1.9254 | -92.1448
1.9291 | -92.3819
1.9328 | -92.4966
1.9365 | -92.4145
1.9402 | -92.2213
1.9438 | -92.0734
1.9475 | -92.085
1.9512 | -92.2695
1.9549 | -92.5569
1.9586 | -92.8288
1.9623 | -92.9681
1.966 | -92.8976
1.9697 | -92.6363
1.9734 | -92.3104
1.977 | -92.0597
1.9807 | -91.9697
1.9844 | -92.0711
1.9881 | -92.3406
1.9918 | -92.6916
1.9955 | -92.9724
1.9992 | -93.0803
2.0029 | -93.0314
2.0065 | -92.913
2.0102 | -92.8581
2.0139 | -92.929
2.0176 | -93.0546
2.0213 | -93.0615
2.025 | -92.8615
2.0287 | -92.5817
2.0324 | -92.4387
2.0361 | -92.5228
2.0397 | -92.7353
2.0434 | -92.8855
2.0471 | -92.8668
2.0508 | -92.7425
2.0545 | -92.6648
2.0582 | -92.7386
2.0619 | -92.9732
2.0656 | -93.302
2.0693 | -93.664
2.0729 | -94.0053
2.0766 | -94.2649
2.0803 | -94.3871
2.084 | -94.3727
2.0877 | -94.3131
2.0914 | -94.2462
2.0951 | -94.0908
2.0988 | -93.7477
2.1024 | -93.2879
2.1061 | -92.8648
2.1098 | -92.5561
2.1135 | -92.3901
2.1172 | -92.3238
2.1209 | -92.2893
2.1246 | -92.2899
2.1283 | -92.3451
2.132 | -92.3738
2.1356 | -92.4052
2.1393 | -92.4478
2.143 | -92.5418
2.1467 | -92.7382
2.1504 | -93.0171
2.1541 | -93.3286
2.1578 | -93.6049
2.1615 | -93.7618
2.1652 | -93.8034
2.1688 | -93.7678
2.1725 | -93.6909
2.1762 | -93.593
2.1799 | -93.4975
2.1836 | -93.41
2.1873 | -93.3235
2.191 | -93.2616
2.1947 | -93.2302
2.1983 | -93.2305
2.202 | -93.2415
2.2057 | -93.2315
2.2094 | -93.2077
2.2131 | -93.2074
2.2168 | -93.261
2.2205 | -93.3524
2.2242 | -93.4552
2.2279 | -93.5654
2.2315 | -93.6839
2.2352 | -93.7984
2.2389 | -93.876
2.2426 | -93.8945
2.2463 | -93.8713
2.25 | -93.8431
2.2537 | -93.8214
2.2574 | -93.8239
2.2611 | -93.8501
2.2647 | -93.8881
2.2684 | -93.9024
2.2721 | -93.8607
2.2758 | -93.7584
2.2795 | -93.6183
2.2832 | -93.4911
2.2869 | -93.4188
2.2906 | -93.4281
2.2943 | -93.5154
2.2979 | -93.6305
2.3016 | -93.7152
2.3053 | -93.7248
2.309 | -93.6372
2.3127 | -93.4713
2.3164 | -93.2966
2.3201 | -93.1896
2.3238 | -93.1884
2.3274 | -93.2741
2.3311 | -93.4024
2.3348 | -93.5194
2.3385 | -93.6037
2.3422 | -93.6217
2.3459 | -93.5564
2.3496 | -93.4396
2.3533 | -93.3279
2.357 | -93.2799
2.3606 | -93.2953
2.3643 | -93.3233
2.368 | -93.3351
2.3717 | -93.3357
2.3754 | -93.3724
2.3791 | -93.4488
2.3828 | -93.5566
2.3865 | -93.6632
2.3902 | -93.7589
2.3938 | -93.8462
2.3975 | -93.9273
2.4012 | -94.0195
2.4049 | -94.1249
2.4086 | -94.2102
2.4123 | -94.2456
2.416 | -94.2134
2.4197 | -94.164
2.4233 | -94.1572
2.427 | -94.2416
2.4307 | -94.3859
2.4344 | -94.5247
2.4381 | -94.5911
2.4418 | -94.5974
2.4455 | -94.5417
2.4492 | -94.4588
2.4529 | -94.3849
2.4565 | -94.349
2.4602 | -94.3745
2.4639 | -94.4477
2.4676 | -94.5169
2.4713 | -94.5523
2.475 | -94.5407
2.4787 | -94.499
2.4824 | -94.4224
2.4861 | -94.3353
2.4897 | -94.2328
2.4934 | -94.1295
2.4971 | -94.0642
2.5008 | -94.056
2.5045 | -94.0995
2.5082 | -94.1644
2.5119 | -94.2177
2.5156 | -94.2548
2.5192 | -94.2924
2.5229 | -94.3411
2.5266 | -94.4157
2.5303 | -94.5165
2.534 | -94.6393
2.5377 | -94.7272
2.5414 | -94.727
2.5451 | -94.6311
2.5488 | -94.4554
2.5524 | -94.2378
2.5561 | -94.0476
2.5598 | -93.9683
2.5635 | -94.0555
2.5672 | -94.2836
2.5709 | -94.5586
2.5746 | -94.7687
2.5783 | -94.8706
2.582 | -94.8664
2.5856 | -94.7993
2.5893 | -94.747
2.593 | -94.7572
2.5967 | -94.8429
2.6004 | -94.9669
2.6041 | -95.0738
2.6078 | -95.1173
2.6115 | -95.101
2.6152 | -95.0797
2.6188 | -95.1047
2.6225 | -95.2232
2.6262 | -95.4212
2.6299 | -95.6279
2.6336 | -95.7318
2.6373 | -95.6381
2.641 | -95.3764
2.6447 | -95.0294
2.6483 | -94.7074
2.652 | -94.4873
2.6557 | -94.4053
2.6594 | -94.5335
2.6631 | -94.618
2.6668 | -94.7406
2.6705 | -94.8713
2.6742 | -94.9871
2.6779 | -95.0873
2.6815 | -95.1805
2.6852 | -95.2517
2.6889 | -95.3051
2.6926 | -95.328
2.6963 | -95.2932
2.7 | -95.1828
2.7037 | -95.0293
2.7074 | -94.8713
2.7111 | -94.7336
2.7147 | -94.6372
2.7184 | -94.563
2.7221 | -94.5058
2.7258 | -94.4595
2.7295 | -94.4475
2.7332 | -94.4892
2.7369 | -94.587
2.7406 | -94.7244
2.7442 | -94.8942
2.7479 | -95.0642
2.7516 | -95.1967
2.7553 | -95.2884
2.759 | -95.3499
2.7627 | -95.3926
2.7664 | -95.4576
2.7701 | -95.5396
2.7738 | -95.6111
2.7774 | -95.6464
2.7811 | -95.6493
2.7848 | -95.6135
2.7885 | -95.5376
2.7922 | -95.4424
2.7959 | -95.3603
2.7996 | -95.3181
2.8033 | -95.3028
2.807 | -95.2858
2.8106 | -95.2423
2.8143 | -95.1491
2.818 | -95.0155
2.8217 | -94.8807
2.8254 | -94.7858
2.8291 | -94.7662
2.8328 | -94.8427
2.8365 | -95.0037
2.8401 | -95.2149
2.8438 | -95.4064
2.8475 | -95.5195
2.8512 | -95.5528
2.8549 | -95.5332
2.8586 | -95.4862
2.8623 | -95.4437
2.866 | -95.4137
2.8697 | -95.3973
2.8733 | -95.3927
2.877 | -95.4219
2.8807 | -95.4981
2.8844 | -95.6086
2.8881 | -95.726
2.8918 | -95.8233
2.8955 | -95.8482
2.8992 | -95.7935
2.9029 | -95.6974
2.9065 | -95.6216
2.9102 | -95.6185
2.9139 | -95.7152
2.9176 | -95.9079
2.9213 | -96.1594
2.925 | -96.4131
2.9287 | -96.6013
2.9324 | -96.6624
2.9361 | -96.5625
2.9397 | -96.3498
2.9434 | -96.0939
2.9471 | -95.8469
2.9508 | -95.6521
2.9545 | -95.5347
2.9582 | -95.5043
2.9619 | -95.5545
2.9656 | -95.6661
2.9692 | -95.8161
2.9729 | -95.9524
2.9766 | -96.0256
2.9803 | -96.0135
2.984 | -95.9257
2.9877 | -95.7776
2.9914 | -95.6074
2.9951 | -95.4393
2.9988 | -95.3077
3.0024 | -95.2294
3.0061 | -95.1914
3.0098 | -95.1937
3.0135 | -95.2274
3.0172 | -95.2758
3.0209 | -95.3418
3.0246 | -95.432
3.0283 | -95.527
3.032 | -95.6228
3.0356 | -95.706
3.0393 | -95.7605
3.043 | -95.7907
3.0467 | -95.7969
3.0504 | -95.7877
3.0541 | -95.7656
3.0578 | -95.703
3.0615 | -95.6027
3.0651 | -95.4976
3.0688 | -95.399
3.0725 | -95.3209
3.0762 | -95.2955
3.0799 | -95.3266
3.0836 | -95.4273
3.0873 | -95.5895
3.091 | -95.7795
3.0947 | -95.9669
3.0983 | -96.1409
3.102 | -96.3053
3.1057 | -96.473
3.1094 | -96.6356
3.1131 | -96.7594
3.1168 | -96.8546
3.1205 | -96.9162
3.1242 | -96.923
3.1279 | -96.8935
3.1315 | -96.8276
3.1352 | -96.7132
3.1389 | -96.5749
3.1426 | -96.4239
3.1463 | -96.2881
3.15 | -96.1834
3.1537 | -96.1141
3.1574 | -96.0653
3.161 | -96.03
3.1647 | -95.9964
3.1684 | -95.9833
3.1721 | -96.0193
3.1758 | -96.1004
3.1795 | -96.2008
3.1832 | -96.2932
3.1869 | -96.3496
3.1906 | -96.3967
3.1942 | -96.4681
3.1979 | -96.5482
3.2016 | -96.6234
3.2053 | -96.6682
3.209 | -96.6652
3.2127 | -96.5953
3.2164 | -96.4555
3.2201 | -96.2763
3.2238 | -96.1016
3.2274 | -95.9798
3.2311 | -95.9441
3.2348 | -95.9876
3.2385 | -96.0933
3.2422 | -96.2429
3.2459 | -96.4051
3.2496 | -96.5634
3.2533 | -96.7024
3.257 | -96.814
3.2606 | -96.8946
3.2643 | -96.9544
3.268 | -97.0018
3.2717 | -97.0503
3.2754 | -97.0821
3.2791 | -97.0931
3.2828 | -97.077
3.2865 | -97.0435
3.2901 | -96.9996
3.2938 | -96.9462
3.2975 | -96.8865
3.3012 | -96.8393
3.3049 | -96.8074
3.3086 | -96.7856
3.3123 | -96.7655
3.316 | -96.7227
3.3197 | -96.6608
3.3233 | -96.5823
3.327 | -96.4829
3.3307 | -96.3914
3.3344 | -96.3106
3.3381 | -96.2429
3.3418 | -96.1935
3.3455 | -96.1588
3.3492 | -96.128
3.3529 | -96.0964
3.3565 | -96.0759
3.3602 | -96.0986
3.3639 | -96.1794
3.3676 | -96.312
3.3713 | -96.4866
3.375 | -96.6638
3.3787 | -96.818
3.3824 | -96.9297
3.386 | -96.983
3.3897 | -96.9913
3.3934 | -96.9944
3.3971 | -97.0129
3.4008 | -97.0546
3.4045 | -97.1228
3.4082 | -97.2018
3.4119 | -97.2615
3.4156 | -97.2833
3.4192 | -97.2824
3.4229 | -97.2731
3.4266 | -97.2506
3.4303 | -97.2216
3.434 | -97.1955
3.4377 | -97.1719
3.4414 | -97.1483
3.4451 | -97.1375
3.4488 | -97.1281
3.4524 | -97.1161
3.4561 | -97.0993
3.4598 | -97.0785
3.4635 | -97.0708
3.4672 | -97.0801
3.4709 | -97.1107
3.4746 | -97.1415
3.4783 | -97.1538
3.4819 | -97.1377
3.4856 | -97.0783
3.4893 | -96.9823
3.493 | -96.8838
3.4967 | -96.7967
3.5004 | -96.7389
3.5041 | -96.7316
3.5078 | -96.7768
3.5115 | -96.8706
3.5151 | -96.9979
3.5188 | -97.1379
3.5225 | -97.2843
3.5262 | -97.4038
3.5299 | -97.4933
3.5336 | -97.5442
3.5373 | -97.5487
3.541 | -97.5197
3.5447 | -97.4688
3.5483 | -97.4127
3.552 | -97.3695
3.5557 | -97.3384
3.5594 | -97.3156
3.5631 | -97.2953
3.5668 | -97.2598
3.5705 | -97.205
3.5742 | -97.1433
3.5779 | -97.0899
3.5815 | -97.0672
3.5852 | -97.0742
3.5889 | -97.117
3.5926 | -97.1895
3.5963 | -97.2706
3.6 | -97.3485
3.6037 | -97.4045
3.6074 | -97.4394
3.611 | -97.4495
3.6147 | -97.4191
3.6184 | -97.362
3.6221 | -97.2949
3.6258 | -97.213
3.6295 | -97.1337
3.6332 | -97.0851
3.6369 | -97.0777
3.6406 | -97.132
3.6442 | -97.2589
3.6479 | -97.44
3.6516 | -97.6514
3.6553 | -97.857
3.659 | -98.0389
3.6627 | -98.1953
3.6664 | -98.3346
3.6701 | -98.4868
3.6738 | -98.6231
3.6774 | -98.715
3.6811 | -98.7467
3.6848 | -98.6883
3.6885 | -98.5273
3.6922 | -98.3122
3.6959 | -98.064
3.6996 | -97.8202
3.7033 | -97.6211
3.7069 | -97.484
3.7106 | -97.4191
3.7143 | -97.4261
3.718 | -97.4988
3.7217 | -97.6659
3.7254 | -97.7962
3.7291 | -97.9241
3.7328 | -98.0436
3.7365 | -98.1242
3.7401 | -98.1526
3.7438 | -98.1295
3.7475 | -98.0714
3.7512 | -97.9842
3.7549 | -97.8959
3.7586 | -97.8104
3.7623 | -97.7479
3.766 | -97.7093
3.7697 | -97.7006
3.7733 | -97.7122
3.777 | -97.7336
3.7807 | -97.7692
3.7844 | -97.8063
3.7881 | -97.8463
3.7918 | -97.8801
3.7955 | -97.9
3.7992 | -97.9178
3.8028 | -97.9244
3.8065 | -97.9314
3.8102 | -97.9542
3.8139 | -97.9793
3.8176 | -98.0004
3.8213 | -98.0164
3.825 | -98.0076
3.8287 | -97.9792
3.8324 | -97.929
3.836 | -97.865
3.8397 | -97.7983
3.8434 | -97.7284
3.8471 | -97.6739
3.8508 | -97.6461
3.8545 | -97.6523
3.8582 | -97.698
3.8619 | -97.7763
3.8656 | -97.8667
3.8692 | -97.9533
3.8729 | -98.0237
3.8766 | -98.0879
3.8803 | -98.1495
3.884 | -98.2052
3.8877 | -98.2568
3.8914 | -98.3009
3.8951 | -98.3232
3.8988 | -98.3209
3.9024 | -98.2746
3.9061 | -98.192
3.9098 | -98.087
3.9135 | -97.9815
3.9172 | -97.8883
3.9209 | -97.8109
3.9246 | -97.7595
3.9283 | -97.7429
3.9319 | -97.7436
3.9356 | -97.7489
3.9393 | -97.7655
3.943 | -97.7924
3.9467 | -97.8248
3.9504 | -97.8632
3.9541 | -97.904
3.9578 | -97.9522
3.9615 | -98.018
3.9651 | -98.1198
3.9688 | -98.2421
3.9725 | -98.3837
3.9762 | -98.5259
3.9799 | -98.6496
3.9836 | -98.7466
3.9873 | -98.8089
3.991 | -98.8318
3.9947 | -98.8265
3.9983 | -98.8047
4.002 | -98.748
4.0057 | -98.6577
4.0094 | -98.558
4.0131 | -98.4633
4.0168 | -98.3924
4.0205 | -98.3651
4.0242 | -98.3821
4.0278 | -98.4441
4.0315 | -98.5353
4.0352 | -98.6156
4.0389 | -98.6663
4.0426 | -98.6821
4.0463 | -98.663
4.05 | -98.6338
4.0537 | -98.6088
4.0574 | -98.6011
4.061 | -98.6231
4.0647 | -98.6656
4.0684 | -98.7163
4.0721 | -98.7554
4.0758 | -98.7891
4.0795 | -98.8094
4.0832 | -98.8098
4.0869 | -98.7905
4.0906 | -98.7543
4.0942 | -98.7117
4.0979 | -98.69
4.1016 | -98.681
4.1053 | -98.6903
4.109 | -98.7001
4.1127 | -98.7016
4.1164 | -98.7012
4.1201 | -98.6897
4.1237 | -98.6816
4.1274 | -98.6927
4.1311 | -98.7277
4.1348 | -98.7878
4.1385 | -98.8527
4.1422 | -98.925
4.1459 | -98.9963
4.1496 | -99.0365
4.1533 | -99.0324
4.1569 | -98.9958
4.1606 | -98.9073
4.1643 | -98.805
4.168 | -98.7028
4.1717 | -98.6207
4.1754 | -98.5683
4.1791 | -98.5532
4.1828 | -98.5704
4.1865 | -98.6036
4.1901 | -98.6457
4.1938 | -98.679
4.1975 | -98.6907
4.2012 | -98.6909
4.2049 | -98.6799
4.2086 | -98.6625
4.2123 | -98.6638
4.216 | -98.6779
4.2197 | -98.7152
4.2233 | -98.7785
4.227 | -98.8501
4.2307 | -98.9246
4.2344 | -98.9813
4.2381 | -99.0063
4.2418 | -98.9895
4.2455 | -98.9474
4.2492 | -98.895
4.2528 | -98.8358
4.2565 | -98.8085
4.2602 | -98.774
4.2639 | -98.7558
4.2676 | -98.7438
4.2713 | -98.7503
4.275 | -98.769
4.2787 | -98.7863
4.2824 | -98.8054
4.286 | -98.8073
4.2897 | -98.7999
4.2934 | -98.7915
4.2971 | -98.7738
4.3008 | -98.7585
4.3045 | -98.749
4.3082 | -98.7572
4.3119 | -98.7829
4.3156 | -98.8301
4.3192 | -98.9156
4.3229 | -99.0192
4.3266 | -99.1381
4.3303 | -99.2579
4.334 | -99.3449
4.3377 | -99.3917
4.3414 | -99.3855
4.3451 | -99.3437
4.3487 | -99.2799
4.3524 | -99.2044
4.3561 | -99.1346
4.3598 | -99.0812
4.3635 | -99.0466
4.3672 | -99.0427
4.3709 | -99.0623
4.3746 | -99.1004
4.3783 | -99.1448
4.3819 | -99.1957
4.3856 | -99.2445
4.3893 | -99.2757
4.393 | -99.2995
4.3967 | -99.3158
4.4004 | -99.3257
4.4041 | -99.3377
4.4078 | -99.3584
4.4115 | -99.3885
4.4151 | -99.4209
4.4188 | -99.4552
4.4225 | -99.4846
4.4262 | -99.5006
4.4299 | -99.4994
4.4336 | -99.4837
4.4373 | -99.4667
4.441 | -99.4538
4.4446 | -99.4492
4.4483 | -99.46
4.452 | -99.4721
4.4557 | -99.4926
4.4594 | -99.5226
4.4631 | -99.5447
4.4668 | -99.5622
4.4705 | -99.571
4.4742 | -99.5713
4.4778 | -99.5572
4.4815 | -99.5253
4.4852 | -99.4776
4.4889 | -99.4134
4.4926 | -99.3549
4.4963 | -99.3064
4.5 | -99.2715
4.5037 | -99.2605
4.5074 | -99.2626
4.511 | -99.2747
4.5147 | -99.2983
4.5184 | -99.3285
4.5221 | -99.3593
4.5258 | -99.3859
4.5295 | -99.4024
4.5332 | -99.4179
4.5369 | -99.4309
4.5406 | -99.4527
4.5442 | -99.4869
4.5479 | -99.5317
4.5516 | -99.5913
4.5553 | -99.6528
4.559 | -99.7068
4.5627 | -99.7335
4.5664 | -99.7391
4.5701 | -99.7189
4.5737 | -99.6749
4.5774 | -99.6268
4.5811 | -99.5832
4.5848 | -99.5397
4.5885 | -99.5069
4.5922 | -99.5011
4.5959 | -99.5109
4.5996 | -99.5345
4.6033 | -99.5566
4.6069 | -99.5795
4.6106 | -99.5948
4.6143 | -99.6046
4.618 | -99.6234
4.6217 | -99.6488
4.6254 | -99.6882
4.6291 | -99.7369
4.6328 | -99.7947
4.6365 | -99.8537
4.6401 | -99.9095
4.6438 | -99.951
4.6475 | -99.9678
4.6512 | -99.97
4.6549 | -99.9522
4.6586 | -99.9136
4.6623 | -99.8567
4.666 | -99.7953
4.6696 | -99.7403
4.6733 | -99.6941
4.677 | -99.6548
4.6807 | -99.64
4.6844 | -99.6511
4.6881 | -99.6711
4.6918 | -99.6905
4.6955 | -99.7109
4.6992 | -99.7232
4.7028 | -99.7131
4.7065 | -99.6799
4.7102 | -99.6226
4.7139 | -99.5507
4.7176 | -99.4748
4.7213 | -99.388
4.725 | -99.3027
4.7287 | -99.2262
4.7324 | -99.1591
4.736 | -99.1054
4.7397 | -99.0716
4.7434 | -99.0581
4.7471 | -99.0531
4.7508 | -99.0595
4.7545 | -99.0749
4.7582 | -99.1046
4.7619 | -99.1382
4.7655 | -99.1634
4.7692 | -99.1886
4.7729 | -99.2171
4.7766 | -99.2405
4.7803 | -99.2726
4.784 | -99.31
4.7877 | -99.3814
4.7914 | -99.4318
4.7951 | -99.4983
4.7987 | -99.5728
4.8024 | -99.6384
4.8061 | -99.6909
4.8098 | -99.7276
4.8135 | -99.739
4.8172 | -99.7259
4.8209 | -99.6948
4.8246 | -99.6392
4.8283 | -99.568
4.8319 | -99.4956
4.8356 | -99.4464
4.8393 | -99.4195
4.843 | -99.4165
4.8467 | -99.4429
4.8504 | -99.4966
4.8541 | -99.563
4.8578 | -99.6452
4.8615 | -99.7394
4.8651 | -99.8432
4.8688 | -99.9548
4.8725 | -100.0654
4.8762 | -100.1711
4.8799 | -100.2662
4.8836 | -100.3434
4.8873 | -100.3856
4.891 | -100.4009
4.8946 | -100.3925
4.8983 | -100.3661
4.902 | -100.3349
4.9057 | -100.3059
4.9094 | -100.2818
4.9131 | -100.2565
4.9168 | -100.2317
4.9205 | -100.2032
4.9242 | -100.1672
4.9278 | -100.1337
4.9315 | -100.0983
4.9352 | -100.048
4.9389 | -99.9915
4.9426 | -99.9311
4.9463 | -99.8676
4.95 | -99.8014
4.9537 | -99.7528
4.9574 | -99.7253
4.961 | -99.7118
4.9647 | -99.7193
4.9684 | -99.741
4.9721 | -99.772
4.9758 | -99.8063
4.9795 | -99.8338
4.9832 | -99.8515
4.9869 | -99.8667
4.9905 | -99.8775
4.9942 | -99.8975
4.9979 | -99.9273
5.0016 | -99.9626
5.0053 | -100.0052
5.009 | -100.052
5.0127 | -100.1113
5.0164 | -100.1609
5.0201 | -100.2035
5.0237 | -100.2365
5.0274 | -100.2598
5.0311 | -100.2558
5.0348 | -100.2395
5.0385 | -100.2099
5.0422 | -100.1573
5.0459 | -100.0914
5.0496 | -100.021
5.0533 | -99.946
5.0569 | -99.8711
5.0606 | -99.8068
5.0643 | -99.745
5.068 | -99.6971
5.0717 | -99.6653
5.0754 | -99.6476
5.0791 | -99.6445
5.0828 | -99.6531
5.0864 | -99.6774
5.0901 | -99.717
5.0938 | -99.7709
5.0975 | -99.8291
5.1012 | -99.8985
5.1049 | -99.9649
5.1086 | -100.0356
5.1123 | -100.0971
5.116 | -100.1572
5.1196 | -100.2159
5.1233 | -100.2636
5.127 | -100.3079
5.1307 | -100.3615
5.1344 | -100.4245
5.1381 | -100.4805
5.1418 | -100.5346
5.1455 | -100.5855
5.1492 | -100.6277
5.1528 | -100.6597
5.1565 | -100.6669
5.1602 | -100.6428
5.1639 | -100.5966
5.1676 | -100.5305
5.1713 | -100.4623
5.175 | -100.401
5.1787 | -100.3503
5.1824 | -100.3302
5.186 | -100.3252
5.1897 | -100.3432
5.1934 | -100.3773
5.1971 | -100.4216
5.2008 | -100.469
5.2045 | -100.5108
5.2082 | -100.5395
5.2119 | -100.5554
5.2155 | -100.5662
5.2192 | -100.5678
5.2229 | -100.5589
5.2266 | -100.5382
5.2303 | -100.5159
5.234 | -100.5014
5.2377 | -100.4914
5.2414 | -100.4792
5.2451 | -100.4685
5.2487 | -100.4544
5.2524 | -100.4365
5.2561 | -100.4228
5.2598 | -100.4073
5.2635 | -100.4027
5.2672 | -100.403
5.2709 | -100.4148
5.2746 | -100.4423
5.2783 | -100.479
5.2819 | -100.5215
5.2856 | -100.5707
5.2893 | -100.6211
5.293 | -100.6688
5.2967 | -100.7026
5.3004 | -100.7227
5.3041 | -100.7299
5.3078 | -100.7249
5.3114 | -100.7077
5.3151 | -100.6829
5.3188 | -100.6382
5.3225 | -100.6032
5.3262 | -100.5657
5.3299 | -100.5316
5.3336 | -100.5032
5.3373 | -100.473
5.341 | -100.4608
5.3446 | -100.4639
5.3483 | -100.4745
5.352 | -100.4937
5.3557 | -100.5194
5.3594 | -100.5546
5.3631 | -100.5878
5.3668 | -100.6096
5.3705 | -100.6156
5.3742 | -100.6099
5.3778 | -100.5942
5.3815 | -100.5511
5.3852 | -100.4897
5.3889 | -100.4153
5.3926 | -100.3354
5.3963 | -100.2627
5.4 | -100.2011
5.4037 | -100.153
5.4073 | -100.1206
5.411 | -100.1035
5.4147 | -100.1023
5.4184 | -100.103
5.4221 | -100.1095
5.4258 | -100.1068
5.4295 | -100.0974
5.4332 | -100.0796
5.4369 | -100.0534
5.4405 | -100.0271
5.4442 | -100.0008
5.4479 | -99.9828
5.4516 | -99.9819
5.4553 | -100.0012
5.459 | -100.0333
5.4627 | -100.0776
5.4664 | -100.1443
5.4701 | -100.2209
5.4737 | -100.299
5.4774 | -100.3801
5.4811 | -100.4536
5.4848 | -100.5197
5.4885 | -100.5743
5.4922 | -100.6179
5.4959 | -100.6568
5.4996 | -100.6845
5.5033 | -100.7127
5.5069 | -100.7439
5.5106 | -100.7873
5.5143 | -100.8326
5.518 | -100.8833
5.5217 | -100.9347
5.5254 | -100.9787
5.5291 | -101.0129
5.5328 | -101.0317
5.5364 | -101.0298
5.5401 | -101.0091
5.5438 | -100.9805
5.5475 | -100.9426
5.5512 | -100.887
5.5549 | -100.8234
5.5586 | -100.7588
5.5623 | -100.693
5.566 | -100.6395
5.5696 | -100.6036
5.5733 | -100.5797
5.577 | -100.5701
5.5807 | -100.5707
5.5844 | -100.5856
5.5881 | -100.6046
5.5918 | -100.6302
5.5955 | -100.6534
5.5992 | -100.6854
5.6028 | -100.7303
5.6065 | -100.7739
5.6102 | -100.8185
5.6139 | -100.8617
5.6176 | -100.8995
5.6213 | -100.933
5.625 | -100.9623
5.6287 | -100.9794
5.6323 | -100.9962
5.636 | -101.0094
5.6397 | -101.0184
5.6434 | -101.0221
5.6471 | -101.0195
5.6508 | -101.0077
5.6545 | -100.9861
5.6582 | -100.9601
5.6619 | -100.9301
5.6655 | -100.9083
5.6692 | -100.8982
5.6729 | -100.9059
5.6766 | -100.9221
5.6803 | -100.954
5.684 | -100.9982
5.6877 | -101.0491
5.6914 | -101.1138
5.6951 | -101.1868
5.6987 | -101.2708
5.7024 | -101.355
5.7061 | -101.4379
5.7098 | -101.5104
5.7135 | -101.5704
5.7172 | -101.6041
5.7209 | -101.6186
5.7246 | -101.61
5.7282 | -101.5727
5.7319 | -101.515
5.7356 | -101.4368
5.7393 | -101.3555
5.743 | -101.2766
5.7467 | -101.2089
5.7504 | -101.151
5.7541 | -101.1141
5.7578 | -101.0931
5.7614 | -101.0951
5.7651 | -101.1198
5.7688 | -101.1587
5.7725 | -101.2064
5.7762 | -101.26
5.7799 | -101.3091
5.7836 | -101.3553
5.7873 | -101.4029
5.791 | -101.4427
5.7946 | -101.4776
5.7983 | -101.5063
5.802 | -101.5335
5.8057 | -101.5462
5.8094 | -101.5486
5.8131 | -101.5451
5.8168 | -101.5311
5.8205 | -101.5108
5.8242 | -101.4817
5.8278 | -101.448
5.8315 | -101.4103
5.8352 | -101.3673
5.8389 | -101.3196
5.8426 | -101.2757
5.8463 | -101.2249
5.85 | -101.1798
5.8537 | -101.1373
5.8573 | -101.1017
5.861 | -101.0773
5.8647 | -101.0645
5.8684 | -101.0669
5.8721 | -101.0785
5.8758 | -101.1054
5.8795 | -101.1404
5.8832 | -101.1784
5.8869 | -101.2211
5.8905 | -101.2579
5.8942 | -101.2898
5.8979 | -101.3056
5.9016 | -101.3065
5.9053 | -101.2967
5.909 | -101.2767
5.9127 | -101.24
5.9164 | -101.205
5.9201 | -101.1683
5.9237 | -101.1393
5.9274 | -101.122
5.9311 | -101.1171
5.9348 | -101.1303
5.9385 | -101.1606
5.9422 | -101.2062
5.9459 | -101.2653
5.9496 | -101.3343
5.9532 | -101.4047
5.9569 | -101.4789
5.9606 | -101.5567
5.9643 | -101.6327
5.968 | -101.6959
5.9717 | -101.7527
5.9754 | -101.8
5.9791 | -101.8195
5.9828 | -101.8172
5.9864 | -101.7896
5.9901 | -101.7421
5.9938 | -101.6786
5.9975 | -101.598
6.0012 | -101.5072
6.0049 | -101.42
6.0086 | -101.325
6.0123 | -101.2434
6.016 | -101.18
6.0196 | -101.1273
6.0233 | -101.0821
6.027 | -101.0467
6.0307 | -101.0229
6.0344 | -101.0062
6.0381 | -100.9933
6.0418 | -100.9885
6.0455 | -100.9975
6.0491 | -101.013
6.0528 | -101.0352
6.0565 | -101.0695
6.0602 | -101.1182
6.0639 | -101.1742
6.0676 | -101.232
6.0713 | -101.2901
6.075 | -101.3509
6.0787 | -101.4132
6.0823 | -101.4638
6.086 | -101.5052
6.0897 | -101.5395
6.0934 | -101.567
6.0971 | -101.5835
6.1008 | -101.5977
6.1045 | -101.6067
6.1082 | -101.6101
6.1119 | -101.6213
6.1155 | -101.6463
6.1192 | -101.6793
6.1229 | -101.7165
6.1266 | -101.7543
6.1303 | -101.7945
6.134 | -101.828
6.1377 | -101.8534
6.1414 | -101.8759
6.1451 | -101.8852
6.1487 | -101.8744
6.1524 | -101.8551
6.1561 | -101.8277
6.1598 | -101.7915
6.1635 | -101.7478
6.1672 | -101.6854
6.1709 | -101.6237
6.1746 | -101.5754
6.1782 | -101.5283
6.1819 | -101.4908
6.1856 | -101.4679
6.1893 | -101.4574
6.193 | -101.459
6.1967 | -101.465
6.2004 | -101.4776
6.2041 | -101.4867
6.2078 | -101.4878
6.2114 | -101.487
6.2151 | -101.4779
6.2188 | -101.4629
6.2225 | -101.4377
6.2262 | -101.4083
6.2299 | -101.3789
6.2336 | -101.3496
6.2373 | -101.3318
6.241 | -101.3243
6.2446 | -101.3138
6.2483 | -101.31
6.252 | -101.3099
6.2557 | -101.3259
6.2594 | -101.3522
6.2631 | -101.3767
6.2668 | -101.407
6.2705 | -101.4443
6.2741 | -101.4813
6.2778 | -101.5188
6.2815 | -101.5595
6.2852 | -101.5981
6.2889 | -101.6371
6.2926 | -101.6698
6.2963 | -101.6986
6.3 | -101.7306
6.3037 | -101.765
6.3073 | -101.8053
6.311 | -101.8445
6.3147 | -101.8863
6.3184 | -101.9327
6.3221 | -101.9773
6.3258 | -102.0207
6.3295 | -102.0603
6.3332 | -102.0997
6.3369 | -102.142
6.3405 | -102.1852
6.3442 | -102.2243
6.3479 | -102.2479
6.3516 | -102.2638
6.3553 | -102.2711
6.359 | -102.2625
6.3627 | -102.2445
6.3664 | -102.2192
6.37 | -102.1834
6.3737 | -102.1419
6.3774 | -102.1046
6.3811 | -102.0702
6.3848 | -102.0602
6.3885 | -102.0464
6.3922 | -102.0445
6.3959 | -102.0448
6.3996 | -102.0464
6.4032 | -102.0543
6.4069 | -102.0706
6.4106 | -102.0859
6.4143 | -102.1
6.418 | -102.1171
6.4217 | -102.1325
6.4254 | -102.1406
6.4291 | -102.1467
6.4328 | -102.1503
6.4364 | -102.143
6.4401 | -102.1366
6.4438 | -102.1287
6.4475 | -102.119
6.4512 | -102.113
6.4549 | -102.1107
6.4586 | -102.109
6.4623 | -102.1127
6.466 | -102.1218
6.4696 | -102.1296
6.4733 | -102.1363
6.477 | -102.1496
6.4807 | -102.167
6.4844 | -102.1771
6.4881 | -102.182
6.4918 | -102.1817
6.4955 | -102.1782
6.4991 | -102.1671
6.5028 | -102.1445
6.5065 | -102.1241
6.5102 | -102.0971
6.5139 | -102.0633
6.5176 | -102.0309
6.5213 | -102.0035
6.525 | -101.9834
6.5287 | -101.9632
6.5323 | -101.9467
6.536 | -101.936
6.5397 | -101.9295
6.5434 | -101.9246
6.5471 | -101.9267
6.5508 | -101.9307
6.5545 | -101.937
6.5582 | -101.9476
6.5619 | -101.955
6.5655 | -101.9633
6.5692 | -101.9777
6.5729 | -101.9936
6.5766 | -102.0018
6.5803 | -102.0095
6.584 | -102.0153
6.5877 | -102.0221
6.5914 | -102.0314
6.595 | -102.0359
6.5987 | -102.039
6.6024 | -102.0489
6.6061 | -102.0546
6.6098 | -102.066
6.6135 | -102.0789
6.6172 | -102.0923
6.6209 | -102.1028
6.6246 | -102.1077
6.6282 | -102.112
6.6319 | -102.1199
6.6356 | -102.1263
6.6393 | -102.1364
6.643 | -102.1531
6.6467 | -102.1687
6.6504 | -102.1882
6.6541 | -102.2088
6.6578 | -102.2275
6.6614 | -102.2386
6.6651 | -102.2498
6.6688 | -102.2555
6.6725 | -102.2556
6.6762 | -102.2571
6.6799 | -102.2511
6.6836 | -102.2488
6.6873 | -102.2514
6.6909 | -102.2556
6.6946 | -102.2643
6.6983 | -102.2787
6.702 | -102.2931
6.7057 | -102.3123
6.7094 | -102.3322
6.7131 | -102.3559
6.7168 | -102.3778
6.7205 | -102.406
6.7241 | -102.4389
6.7278 | -102.4664
6.7315 | -102.4827
6.7352 | -102.4998
6.7389 | -102.4987
6.7426 | -102.4876
6.7463 | -102.4699
6.75 | -102.441
6.7537 | -102.4064
6.7573 | -102.3742
6.761 | -102.3417
6.7647 | -102.3044
6.7684 | -102.2743
6.7721 | -102.252
6.7758 | -102.2374
6.7795 | -102.2381
6.7832 | -102.2417
6.7869 | -102.2521
6.7905 | -102.2701
6.7942 | -102.2983
6.7979 | -102.3233
6.8016 | -102.3434
6.8053 | -102.3556
6.809 | -102.3656
6.8127 | -102.3683
6.8164 | -102.3583
6.82 | -102.3393
6.8237 | -102.3226
6.8274 | -102.3059
6.8311 | -102.2836
6.8348 | -102.2626
6.8385 | -102.2461
6.8422 | -102.2307
6.8459 | -102.218
6.8496 | -102.2061
6.8532 | -102.1977
6.8569 | -102.1902
6.8606 | -102.1846
6.8643 | -102.1815
6.868 | -102.1817
6.8717 | -102.1848
6.8754 | -102.1952
6.8791 | -102.2054
6.8828 | -102.2212
6.8864 | -102.2341
6.8901 | -102.241
6.8938 | -102.2573
6.8975 | -102.2831
6.9012 | -102.3129
6.9049 | -102.3447
6.9086 | -102.3789
6.9123 | -102.4192
6.9159 | -102.4637
6.9196 | -102.5033
6.9233 | -102.5393
6.927 | -102.5689
6.9307 | -102.5978
6.9344 | -102.6241
6.9381 | -102.6454
6.9418 | -102.6531
6.9455 | -102.6575
6.9491 | -102.6551
6.9528 | -102.647
6.9565 | -102.6279
6.9602 | -102.6003
6.9639 | -102.5716
6.9676 | -102.5361
6.9713 | -102.4925
6.975 | -102.4513
6.9787 | -102.4133
6.9823 | -102.384
6.986 | -102.3662
6.9897 | -102.3612
6.9934 | -102.3694
6.9971 | -102.3764
7.0008 | -102.3891
7.0045 | -102.4163
7.0082 | -102.454
7.0118 | -102.4917
7.0155 | -102.5278
7.0192 | -102.5621
7.0229 | -102.5959
7.0266 | -102.6263
7.0303 | -102.6458
7.034 | -102.6612
7.0377 | -102.6692
7.0414 | -102.6723
7.045 | -102.6615
7.0487 | -102.6494
7.0524 | -102.6291
7.0561 | -102.6123
7.0598 | -102.5977
7.0635 | -102.5873
7.0672 | -102.5838
7.0709 | -102.583
7.0746 | -102.5873
7.0782 | -102.5999
7.0819 | -102.6111
7.0856 | -102.6174
7.0893 | -102.6262
7.093 | -102.6387
7.0967 | -102.6505
7.1004 | -102.6616
7.1041 | -102.6758
7.1078 | -102.6881
7.1114 | -102.6942
7.1151 | -102.7013
7.1188 | -102.7134
7.1225 | -102.727
7.1262 | -102.744
7.1299 | -102.7609
7.1336 | -102.7786
7.1373 | -102.7961
7.1409 | -102.8036
7.1446 | -102.8188
7.1483 | -102.8383
7.152 | -102.8547
7.1557 | -102.86
7.1594 | -102.8565
7.1631 | -102.8489
7.1668 | -102.8332
7.1705 | -102.8105
7.1741 | -102.7834
7.1778 | -102.7595
7.1815 | -102.7328
7.1852 | -102.7068
7.1889 | -102.6918
7.1926 | -102.686
7.1963 | -102.6865
7.2 | -102.6875
7.2037 | -102.6931
7.2073 | -102.7096
7.211 | -102.7315
7.2147 | -102.7588
7.2184 | -102.7881
7.2221 | -102.8144
7.2258 | -102.8372
7.2295 | -102.8623
7.2332 | -102.8855
7.2368 | -102.9072
7.2405 | -102.9187
7.2442 | -102.922
7.2479 | -102.9217
7.2516 | -102.9104
7.2553 | -102.8928
7.259 | -102.8741
7.2627 | -102.8568
7.2664 | -102.8344
7.27 | -102.8096
7.2737 | -102.7875
7.2774 | -102.7668
7.2811 | -102.7482
7.2848 | -102.7316
7.2885 | -102.7261
7.2922 | -102.73
7.2959 | -102.7351
7.2996 | -102.7453
7.3032 | -102.7553
7.3069 | -102.7716
7.3106 | -102.793
7.3143 | -102.8123
7.318 | -102.8284
7.3217 | -102.8354
7.3254 | -102.8369
7.3291 | -102.8426
7.3327 | -102.8438
7.3364 | -102.8336
7.3401 | -102.8108
7.3438 | -102.7856
7.3475 | -102.7633
7.3512 | -102.7446
7.3549 | -102.7232
7.3586 | -102.6972
7.3623 | -102.6744
7.3659 | -102.6577
7.3696 | -102.6451
7.3733 | -102.6434
7.377 | -102.6465
7.3807 | -102.6574
7.3844 | -102.6743
7.3881 | -102.6933
7.3918 | -102.7122
7.3955 | -102.7334
7.3991 | -102.7526
7.4028 | -102.7749
7.4065 | -102.7996
7.4102 | -102.8234
7.4139 | -102.8383
7.4176 | -102.8466
7.4213 | -102.8542
7.425 | -102.8535
7.4287 | -102.8523
7.4323 | -102.8509
7.436 | -102.8329
7.4397 | -102.8182
7.4434 | -102.8066
7.4471 | -102.7985
7.4508 | -102.7938
7.4545 | -102.787
7.4582 | -102.7789
7.4618 | -102.7716
7.4655 | -102.774
7.4692 | -102.7809
7.4729 | -102.7887
7.4766 | -102.7986
7.4803 | -102.8116
7.484 | -102.818
7.4877 | -102.8246
7.4914 | -102.8365
7.495 | -102.8486
7.4987 | -102.8571
7.5024 | -102.8636
7.5061 | -102.861
7.5098 | -102.8548
7.5135 | -102.8523
7.5172 | -102.844
7.5209 | -102.839
7.5246 | -102.8383
7.5282 | -102.8316
7.5319 | -102.8254
7.5356 | -102.8316
7.5393 | -102.8411
7.543 | -102.8445
7.5467 | -102.8529
7.5504 | -102.8638
7.5541 | -102.8799
7.5577 | -102.8992
7.5614 | -102.9147
7.5651 | -102.9371
7.5688 | -102.9622
7.5725 | -102.9828
7.5762 | -103.0044
7.5799 | -103.0216
7.5836 | -103.0393
7.5873 | -103.0521
7.5909 | -103.0609
7.5946 | -103.0678
7.5983 | -103.0796
7.602 | -103.0877
7.6057 | -103.0874
7.6094 | -103.0866
7.6131 | -103.0862
7.6168 | -103.0868
7.6205 | -103.0954
7.6241 | -103.1017
7.6278 | -103.1009
7.6315 | -103.111
7.6352 | -103.1217
7.6389 | -103.1342
7.6426 | -103.1519
7.6463 | -103.1684
7.65 | -103.1806
7.6536 | -103.1948
7.6573 | -103.2144
7.661 | -103.2396
7.6647 | -103.2641
7.6684 | -103.2846
7.6721 | -103.3042
7.6758 | -103.3278
7.6795 | -103.3459
7.6832 | -103.364
7.6868 | -103.3844
7.6905 | -103.3993
7.6942 | -103.4108
7.6979 | -103.4242
7.7016 | -103.4342
7.7053 | -103.4389
7.709 | -103.4388
7.7127 | -103.4293
7.7164 | -103.4219
7.72 | -103.4185
7.7237 | -103.4127
7.7274 | -103.4049
7.7311 | -103.3929
7.7348 | -103.3735
7.7385 | -103.3515
7.7422 | -103.3278
7.7459 | -103.2993
7.7496 | -103.2704
7.7532 | -103.2344
7.7569 | -103.2016
7.7606 | -103.1727
7.7643 | -103.1439
7.768 | -103.1198
7.7717 | -103.0974
7.7754 | -103.072
7.7791 | -103.0461
7.7827 | -103.0289
7.7864 | -103.0217
7.7901 | -103.0133
7.7938 | -103.0051
7.7975 | -103.0064
7.8012 | -103.0156
7.8049 | -103.0279
7.8086 | -103.0359
7.8123 | -103.0443
7.8159 | -103.0624
7.8196 | -103.0816
7.8233 | -103.0968
7.827 | -103.1123
7.8307 | -103.1155
7.8344 | -103.1167
7.8381 | -103.1189
7.8418 | -103.12
7.8455 | -103.1229
7.8491 | -103.1295
7.8528 | -103.1371
7.8565 | -103.1517
7.8602 | -103.1707
7.8639 | -103.1962
7.8676 | -103.2334
7.8713 | -103.271
7.875 | -103.3052
7.8786 | -103.3453
7.8823 | -103.387
7.886 | -103.4249
7.8897 | -103.4535
7.8934 | -103.4747
7.8971 | -103.4884
7.9008 | -103.4942
7.9045 | -103.4946
7.9082 | -103.4876
7.9118 | -103.4673
7.9155 | -103.4433
7.9192 | -103.4223
7.9229 | -103.4122
7.9266 | -103.3998
7.9303 | -103.3885
7.934 | -103.3842
7.9377 | -103.3901
7.9414 | -103.4091
7.945 | -103.4431
7.9487 | -103.4741
7.9524 | -103.5077
7.9561 | -103.5464
7.9598 | -103.5869
7.9635 | -103.6283
7.9672 | -103.6712
7.9709 | -103.704
7.9745 | -103.7294
7.9782 | -103.7424
7.9819 | -103.7113
7.9856 | -103.705
7.9893 | -103.6909
7.993 | -103.6681
7.9967 | -103.6322
8.0004 | -103.5932
8.0041 | -103.5553
8.0077 | -103.5207
8.0114 | -103.4875
8.0151 | -103.4548
8.0188 | -103.4226
8.0225 | -103.3976
8.0262 | -103.3751
8.0299 | -103.3543
8.0336 | -103.3333
8.0373 | -103.3136
8.0409 | -103.2934
8.0446 | -103.2777
8.0483 | -103.2625
8.052 | -103.2514
8.0557 | -103.2413
8.0594 | -103.2323
8.0631 | -103.2273
8.0668 | -103.2354
8.0704 | -103.2506
8.0741 | -103.2695
8.0778 | -103.2906
8.0815 | -103.3139
8.0852 | -103.3405
8.0889 | -103.3687
8.0926 | -103.3998
8.0963 | -103.4366
8.1 | -103.4781
8.1036 | -103.5129
8.1073 | -103.546
8.111 | -103.5757
8.1147 | -103.5999
8.1184 | -103.6271
8.1221 | -103.6449
8.1258 | -103.6528
8.1295 | -103.6548
8.1332 | -103.6495
8.1368 | -103.6363
8.1405 | -103.6206
8.1442 | -103.6
8.1479 | -103.5812
8.1516 | -103.5628
8.1553 | -103.5475
8.159 | -103.5323
8.1627 | -103.5151
8.1664 | -103.4964
8.17 | -103.4849
8.1737 | -103.4783
8.1774 | -103.4797
8.1811 | -103.4851
8.1848 | -103.493
8.1885 | -103.4984
8.1922 | -103.5055
8.1959 | -103.5176
8.1995 | -103.5347
8.2032 | -103.5539
8.2069 | -103.5751
8.2106 | -103.5936
8.2143 | -103.6069
8.218 | -103.6197
8.2217 | -103.6271
8.2254 | -103.634
8.2291 | -103.6395
8.2327 | -103.6401
8.2364 | -103.6394
8.2401 | -103.6315
8.2438 | -103.6204
8.2475 | -103.6093
8.2512 | -103.5903
8.2549 | -103.5695
8.2586 | -103.5528
8.2623 | -103.5395
8.2659 | -103.528
8.2696 | -103.5157
8.2733 | -103.504
8.277 | -103.4933
8.2807 | -103.4907
8.2844 | -103.493
8.2881 | -103.4998
8.2918 | -103.5082
8.2954 | -103.5207
8.2991 | -103.5362
8.3028 | -103.5599
8.3065 | -103.5884
8.3102 | -103.6197
8.3139 | -103.6536
8.3176 | -103.6908
8.3213 | -103.7242
8.325 | -103.7541
8.3286 | -103.7814
8.3323 | -103.8061
8.336 | -103.8282
8.3397 | -103.843
8.3434 | -103.8538
8.3471 | -103.8618
8.3508 | -103.8644
8.3545 | -103.8688
8.3582 | -103.87
8.3618 | -103.8649
8.3655 | -103.8592
8.3692 | -103.857
8.3729 | -103.8636
8.3766 | -103.8661
8.3803 | -103.8587
8.384 | -103.8473
8.3877 | -103.8373
8.3913 | -103.83
8.395 | -103.8219
8.3987 | -103.8056
8.4024 | -103.7844
8.4061 | -103.7679
8.4098 | -103.7513
8.4135 | -103.7291
8.4172 | -103.7059
8.4209 | -103.6821
8.4245 | -103.6553
8.4282 | -103.6313
8.4319 | -103.6121
8.4356 | -103.5955
8.4393 | -103.5805
8.443 | -103.57
8.4467 | -103.5645
8.4504 | -103.5624
8.4541 | -103.5686
8.4577 | -103.5755
8.4614 | -103.5826
8.4651 | -103.5912
8.4688 | -103.6046
8.4725 | -103.6188
8.4762 | -103.6348
8.4799 | -103.6557
8.4836 | -103.679
8.4873 | -103.7018
8.4909 | -103.7214
8.4946 | -103.741
8.4983 | -103.7658
8.502 | -103.7946
8.5057 | -103.8247
8.5094 | -103.8514
8.5131 | -103.8862
8.5168 | -103.9133
8.5204 | -103.9461
8.5241 | -103.9811
8.5278 | -104.0089
8.5315 | -104.0312
8.5352 | -104.0525
8.5389 | -104.0718
8.5426 | -104.0882
8.5463 | -104.1009
8.55 | -104.106
8.5536 | -104.1027
8.5573 | -104.0919
8.561 | -104.0756
8.5647 | -104.0575
8.5684 | -104.0409
8.5721 | -104.0211
8.5758 | -103.9984
8.5795 | -103.973
8.5832 | -103.9538
8.5868 | -103.9417
8.5905 | -103.9351
8.5942 | -103.9276
8.5979 | -103.922
8.6016 | -103.9179
8.6053 | -103.9176
8.609 | -103.9212
8.6127 | -103.9216
8.6163 | -103.9223
8.62 | -103.9189
8.6237 | -103.9086
8.6274 | -103.8965
8.6311 | -103.8829
8.6348 | -103.8673
8.6385 | -103.8399
8.6422 | -103.8045
8.6459 | -103.7697
8.6495 | -103.7292
8.6532 | -103.6903
8.6569 | -103.6509
8.6606 | -103.61
8.6643 | -103.5741
8.668 | -103.5387
8.6717 | -103.5056
8.6754 | -103.4789
8.6791 | -103.4631
8.6827 | -103.4541
8.6864 | -103.4484
8.6901 | -103.4467
8.6938 | -103.4482
8.6975 | -103.4573
8.7012 | -103.4711
8.7049 | -103.4898
8.7086 | -103.5097
8.7122 | -103.5299
8.7159 | -103.5511
8.7196 | -103.5719
8.7233 | -103.5917
8.727 | -103.6072
8.7307 | -103.622
8.7344 | -103.632
8.7381 | -103.6388
8.7418 | -103.6444
8.7454 | -103.6464
8.7491 | -103.6424
8.7528 | -103.6411
8.7565 | -103.6381
8.7602 | -103.6247
8.7639 | -103.6102
8.7676 | -103.5984
8.7713 | -103.5854
8.775 | -103.5657
8.7786 | -103.5468
8.7823 | -103.53
8.786 | -103.5143
8.7897 | -103.5018
8.7934 | -103.4902
8.7971 | -103.4849
8.8008 | -103.4814
8.8045 | -103.4762
8.8082 | -103.4773
8.8118 | -103.4859
8.8155 | -103.4981
8.8192 | -103.5179
8.8229 | -103.5434
8.8266 | -103.5696
8.8303 | -103.6006
8.834 | -103.6347
8.8377 | -103.6716
8.8413 | -103.707
8.845 | -103.7396
8.8487 | -103.7782
8.8524 | -103.8232
8.8561 | -103.8698
8.8598 | -103.9124
8.8635 | -103.9533
8.8672 | -103.9902
8.8709 | -104.0201
8.8745 | -104.0487
8.8782 | -104.0741
8.8819 | -104.0922
8.8856 | -104.1075
8.8893 | -104.1226
8.893 | -104.1367
8.8967 | -104.1483
8.9004 | -104.1582
8.9041 | -104.1623
8.9077 | -104.1645
8.9114 | -104.1632
8.9151 | -104.1635
8.9188 | -104.1614
8.9225 | -104.1594
8.9262 | -104.1592
8.9299 | -104.1643
8.9336 | -104.1755
8.9372 | -104.1806
8.9409 | -104.1856
8.9446 | -104.193
8.9483 | -104.1994
8.952 | -104.2047
8.9557 | -104.2092
8.9594 | -104.2161
8.9631 | -104.2228
8.9668 | -104.2237
8.9704 | -104.2263
8.9741 | -104.2266
8.9778 | -104.2185
8.9815 | -104.2043
8.9852 | -104.1876
8.9889 | -104.1711
8.9926 | -104.1501
8.9963 | -104.1268
9 | -104.108
9.0036 | -104.0859
9.0073 | -104.0634
9.011 | -104.0406
9.0147 | -104.015
9.0184 | -103.9931
9.0221 | -103.9661
9.0258 | -103.9407
9.0295 | -103.9213
9.0331 | -103.9042
9.0368 | -103.8885
9.0405 | -103.8728
9.0442 | -103.8688
9.0479 | -103.8606
9.0516 | -103.8571
9.0553 | -103.8549
9.059 | -103.8537
9.0627 | -103.8553
9.0663 | -103.8632
9.07 | -103.8743
9.0737 | -103.8879
9.0774 | -103.9043
9.0811 | -103.925
9.0848 | -103.9482
9.0885 | -103.9687
9.0922 | -103.9874
9.0959 | -104.0055
9.0995 | -104.0276
9.1032 | -104.0459
9.1069 | -104.0649
9.1106 | -104.0832
9.1143 | -104.0993
9.118 | -104.1174
9.1217 | -104.1302
9.1254 | -104.1385
9.1291 | -104.1452
9.1327 | -104.149
9.1364 | -104.153
9.1401 | -104.1547
9.1438 | -104.1525
9.1475 | -104.1542
9.1512 | -104.1529
9.1549 | -104.148
9.1586 | -104.1438
9.1622 | -104.1377
9.1659 | -104.1299
9.1696 | -104.1226
9.1733 | -104.1166
9.177 | -104.1067
9.1807 | -104.0962
9.1844 | -104.0874
9.1881 | -104.0817
9.1918 | -104.0801
9.1954 | -104.0809
9.1991 | -104.0833
9.2028 | -104.08
9.2065 | -104.0842
9.2102 | -104.0926
9.2139 | -104.1085
9.2176 | -104.1226
9.2213 | -104.1347
9.225 | -104.1533
9.2286 | -104.1743
9.2323 | -104.1917
9.236 | -104.2079
9.2397 | -104.224
9.2434 | -104.2393
9.2471 | -104.2562
9.2508 | -104.2709
9.2545 | -104.2795
9.2581 | -104.2905
9.2618 | -104.2989
9.2655 | -104.3063
9.2692 | -104.3117
9.2729 | -104.3073
9.2766 | -104.3004
9.2803 | -104.2909
9.284 | -104.2756
9.2877 | -104.2556
9.2913 | -104.2352
9.295 | -104.2048
9.2987 | -104.1736
9.3024 | -104.1413
9.3061 | -104.1052
9.3098 | -104.0669
9.3135 | -104.0322
9.3172 | -103.9957
9.3209 | -103.9632
9.3245 | -103.9342
9.3282 | -103.9105
9.3319 | -103.8934
9.3356 | -103.8806
9.3393 | -103.8799
9.343 | -103.882
9.3467 | -103.8852
9.3504 | -103.8892
9.3541 | -103.9019
9.3577 | -103.9212
9.3614 | -103.9415
9.3651 | -103.962
9.3688 | -103.9859
9.3725 | -104.0105
9.3762 | -104.0329
9.3799 | -104.0573
9.3836 | -104.0812
9.3872 | -104.1061
9.3909 | -104.1304
9.3946 | -104.1476
9.3983 | -104.1645
9.402 | -104.1771
9.4057 | -104.1858
9.4094 | -104.2
9.4131 | -104.2137
9.4168 | -104.2246
9.4204 | -104.2375
9.4241 | -104.2535
9.4278 | -104.2713
9.4315 | -104.298
9.4352 | -104.3223
9.4389 | -104.3481
9.4426 | -104.3779
9.4463 | -104.4131
9.45 | -104.4449
9.4536 | -104.4722
9.4573 | -104.5
9.461 | -104.5227
9.4647 | -104.5426
9.4684 | -104.559
9.4721 | -104.5681
9.4758 | -104.5752
9.4795 | -104.5806
9.4831 | -104.5852
9.4868 | -104.5878
9.4905 | -104.5855
9.4942 | -104.5816
9.4979 | -104.5752
9.5016 | -104.5621
9.5053 | -104.5503
9.509 | -104.5422
9.5127 | -104.5326
9.5163 | -104.5155
9.52 | -104.493
9.5237 | -104.4731
9.5274 | -104.4566
9.5311 | -104.4411
9.5348 | -104.42
9.5385 | -104.3951
9.5422 | -104.3683
9.5459 | -104.3419
9.5495 | -104.3173
9.5532 | -104.2917
9.5569 | -104.2644
9.5606 | -104.2388
9.5643 | -104.214
9.568 | -104.1883
9.5717 | -104.1731
9.5754 | -104.1824
9.579 | -104.1799
9.5827 | -104.1843
9.5864 | -104.193
9.5901 | -104.21
9.5938 | -104.2368
9.5975 | -104.2682
9.6012 | -104.3058
9.6049 | -104.3463
9.6086 | -104.3873
9.6122 | -104.434
9.6159 | -104.4795
9.6196 | -104.5249
9.6233 | -104.5703
9.627 | -104.6122
9.6307 | -104.6473
9.6344 | -104.6772
9.6381 | -104.6984
9.6418 | -104.7153
9.6454 | -104.7223
9.6491 | -104.7124
9.6528 | -104.6957
9.6565 | -104.6752
9.6602 | -104.6535
9.6639 | -104.6224
9.6676 | -104.5833
9.6713 | -104.5401
9.675 | -104.5005
9.6786 | -104.4601
9.6823 | -104.413
9.686 | -104.3675
9.6897 | -104.3294
9.6934 | -104.2965
9.6971 | -104.2655
9.7008 | -104.2412
9.7045 | -104.2229
9.7081 | -104.2099
9.7118 | -104.1992
9.7155 | -104.1947
9.7192 | -104.1967
9.7229 | -104.2065
9.7266 | -104.2181
9.7303 | -104.2332
9.734 | -104.2522
9.7377 | -104.272
9.7413 | -104.296
9.745 | -104.3237
9.7487 | -104.3545
9.7524 | -104.3909
9.7561 | -104.4243
9.7598 | -104.4533
9.7635 | -104.484
9.7672 | -104.5126
9.7709 | -104.5364
9.7745 | -104.563
9.7782 | -104.5894
9.7819 | -104.6146
9.7856 | -104.6331
9.7893 | -104.6477
9.793 | -104.661
9.7967 | -104.6692
9.8004 | -104.6701
9.804 | -104.666
9.8077 | -104.658
9.8114 | -104.6468
9.8151 | -104.6362
9.8188 | -104.6186
9.8225 | -104.5977
9.8262 | -104.5819
9.8299 | -104.5616
9.8336 | -104.5324
9.8372 | -104.5061
9.8409 | -104.4812
9.8446 | -104.4575
9.8483 | -104.4353
9.852 | -104.4147
9.8557 | -104.3956
9.8594 | -104.3765
9.8631 | -104.3587
9.8668 | -104.3417
9.8704 | -104.3278
9.8741 | -104.32
9.8778 | -104.3161
9.8815 | -104.3171
9.8852 | -104.3203
9.8889 | -104.3266
9.8926 | -104.3366
9.8963 | -104.3492
9.8999 | -104.3674
9.9036 | -104.3924
9.9073 | -104.4174
9.911 | -104.4477
9.9147 | -104.4782
9.9184 | -104.5048
9.9221 | -104.5297
9.9258 | -104.5526
9.9295 | -104.5746
9.9331 | -104.5957
9.9368 | -104.6098
9.9405 | -104.6224
9.9442 | -104.6296
9.9479 | -104.6285
9.9516 | -104.6277
9.9553 | -104.625
9.959 | -104.6183
9.9627 | -104.6084
9.9663 | -104.598
9.97 | -104.5841
9.9737 | -104.5699
9.9774 | -104.5578
9.9811 | -104.5422
9.9848 | -104.5309
9.9885 | -104.5172
9.9922 | -104.499
9.9959 | -104.4838
9.9995 | -104.4716
10.0032 | -104.4624
10.0069 | -104.4565
10.0106 | -104.4531
10.0143 | -104.4488
10.018 | -104.444
10.0217 | -104.4416
10.0254 | -104.4443
10.029 | -104.4481
10.0327 | -104.4583
10.0364 | -104.4738
10.0401 | -104.4916
10.0438 | -104.5092
10.0475 | -104.5257
10.0512 | -104.5496
10.0549 | -104.5705
10.0586 | -104.5928
10.0622 | -104.6162
10.0659 | -104.6396
10.0696 | -104.6652
10.0733 | -104.6944
10.077 | -104.7192
10.0807 | -104.7466
10.0844 | -104.7728
10.0881 | -104.7994
10.0918 | -104.8301
10.0954 | -104.8547
10.0991 | -104.8799
10.1028 | -104.899
10.1065 | -104.9116
10.1102 | -104.8936
10.1139 | -104.8963
10.1176 | -104.8919
10.1213 | -104.8809
10.1249 | -104.8661
10.1286 | -104.847
10.1323 | -104.818
10.136 | -104.7876
10.1397 | -104.7596
10.1434 | -104.7276
10.1471 | -104.6921
10.1508 | -104.655
10.1545 | -104.6179
10.1581 | -104.581
10.1618 | -104.55
10.1655 | -104.5272
10.1692 | -104.5078
10.1729 | -104.491
10.1766 | -104.4798
10.1803 | -104.4734
10.184 | -104.4728
10.1877 | -104.478
10.1913 | -104.4906
10.195 | -104.5083
10.1987 | -104.5291
10.2024 | -104.5557
10.2061 | -104.5851
10.2098 | -104.6143
10.2135 | -104.644
10.2172 | -104.6733
10.2208 | -104.6985
10.2245 | -104.7268
10.2282 | -104.7563
10.2319 | -104.7831
10.2356 | -104.8072
10.2393 | -104.8226
10.243 | -104.84
10.2467 | -104.8493
10.2504 | -104.8626
10.254 | -104.8792
10.2577 | -104.8901
10.2614 | -104.8955
10.2651 | -104.8987
10.2688 | -104.9026
10.2725 | -104.9052
10.2762 | -104.9066
10.2799 | -104.9101
10.2836 | -104.9143
10.2872 | -104.9187
10.2909 | -104.9218
10.2946 | -104.9267
10.2983 | -104.9329
10.302 | -104.9365
10.3057 | -104.9355
10.3094 | -104.9372
10.3131 | -104.9393
10.3168 | -104.9377
10.3204 | -104.9396
10.3241 | -104.9424
10.3278 | -104.9454
10.3315 | -104.9488
10.3352 | -104.9523
10.3389 | -104.9549
10.3426 | -104.9565
10.3463 | -104.9585
10.3499 | -104.9575
10.3536 | -104.9582
10.3573 | -104.9583
10.361 | -104.9589
10.3647 | -104.956
10.3684 | -104.949
10.3721 | -104.9428
10.3758 | -104.9336
10.3795 | -104.9194
10.3831 | -104.9052
10.3868 | -104.8905
10.3905 | -104.8794
10.3942 | -104.8646
10.3979 | -104.8534
10.4016 | -104.841
10.4053 | -104.8288
10.409 | -104.8205
10.4127 | -104.8108
10.4163 | -104.8024
10.42 | -104.7964
10.4237 | -104.7906
10.4274 | -104.7899
10.4311 | -104.7898
10.4348 | -104.7883
10.4385 | -104.7889
10.4422 | -104.7917
10.4458 | -104.7915
10.4495 | -104.7918
10.4532 | -104.7971
10.4569 | -104.8033
10.4606 | -104.8034
10.4643 | -104.7999
10.468 | -104.7962
10.4717 | -104.7908
10.4754 | -104.7842
10.479 | -104.779
10.4827 | -104.7744
10.4864 | -104.7646
10.4901 | -104.757
10.4938 | -104.7559
10.4975 | -104.7582
10.5012 | -104.763
10.5049 | -104.7699
10.5086 | -104.781
10.5122 | -104.7879
10.5159 | -104.797
10.5196 | -104.8154
10.5233 | -104.8327
10.527 | -104.8503
10.5307 | -104.8693
10.5344 | -104.8853
10.5381 | -104.9064
10.5417 | -104.9257
10.5454 | -104.9441
10.5491 | -104.9597
10.5528 | -104.9719
10.5565 | -104.9825
10.5602 | -104.9962
10.5639 | -105.0096
10.5676 | -105.0209
10.5713 | -105.0292
10.5749 | -105.0395
10.5786 | -105.0452
10.5823 | -105.0481
10.586 | -105.0516
10.5897 | -105.0539
10.5934 | -105.0525
10.5971 | -105.0548
10.6008 | -105.0545
10.6045 | -105.0501
10.6081 | -105.0437
10.6118 | -105.0373
10.6155 | -105.0278
10.6192 | -105.0172
10.6229 | -105.0052
10.6266 | -104.9953
10.6303 | -104.9858
10.634 | -104.9743
10.6377 | -104.9704
10.6413 | -104.9624
10.645 | -104.9553
10.6487 | -104.9423
10.6524 | -104.9292
10.6561 | -104.9172
10.6598 | -104.9075
10.6635 | -104.9001
10.6672 | -104.8978
10.6708 | -104.8964
10.6745 | -104.8899
10.6782 | -104.8861
10.6819 | -104.8838
10.6856 | -104.8853
10.6893 | -104.8886
10.693 | -104.8888
10.6967 | -104.8904
10.7004 | -104.8915
10.704 | -104.8907
10.7077 | -104.8885
10.7114 | -104.8852
10.7151 | -104.8799
10.7188 | -104.8737
10.7225 | -104.8685
10.7262 | -104.8656
10.7299 | -104.8614
10.7336 | -104.8578
10.7372 | -104.8551
10.7409 | -104.855
10.7446 | -104.8511
10.7483 | -104.8513
10.752 | -104.8501
10.7557 | -104.849
10.7594 | -104.8506
10.7631 | -104.8567
10.7667 | -104.8649
10.7704 | -104.8767
10.7741 | -104.8891
10.7778 | -104.9066
10.7815 | -104.9262
10.7852 | -104.9442
10.7889 | -104.962
10.7926 | -104.9844
10.7963 | -105.0132
10.7999 | -105.0455
10.8036 | -105.0793
10.8073 | -105.1111
10.811 | -105.1455
10.8147 | -105.1815
10.8184 | -105.2148
10.8221 | -105.2413
10.8258 | -105.2693
10.8295 | -105.296
10.8331 | -105.3206
10.8368 | -105.3449
10.8405 | -105.3668
10.8442 | -105.3881
10.8479 | -105.4026
10.8516 | -105.4135
10.8553 | -105.4217
10.859 | -105.4278
10.8626 | -105.432
10.8663 | -105.43
10.87 | -105.4218
10.8737 | -105.4049
10.8774 | -105.3848
10.8811 | -105.3624
10.8848 | -105.3333
10.8885 | -105.2973
10.8922 | -105.264
10.8958 | -105.2301
10.8995 | -105.1912
10.9032 | -105.1533
10.9069 | -105.1139
10.9106 | -105.0744
10.9143 | -105.041
10.918 | -105.0052
10.9217 | -104.9663
10.9254 | -104.9337
10.929 | -104.9036
10.9327 | -104.8735
10.9364 | -104.847
10.9401 | -104.8279
10.9438 | -104.8158
10.9475 | -104.8098
10.9512 | -104.8039
10.9549 | -104.802
10.9586 | -104.8063
10.9622 | -104.8178
10.9659 | -104.8355
10.9696 | -104.8527
10.9733 | -104.8727
10.977 | -104.8946
10.9807 | -104.9185
10.9844 | -104.9428
10.9881 | -104.9693
10.9917 | -104.9977
10.9954 | -105.0216
10.9991 | -105.0411
11.0028 | -105.0588
11.0065 | -105.0781
11.0102 | -105.0971
11.0139 | -105.1143
11.0176 | -105.1319
11.0213 | -105.1498
11.0249 | -105.1662
11.0286 | -105.1787
11.0323 | -105.1908
11.036 | -105.2004
11.0397 | -105.2091
11.0434 | -105.2161
11.0471 | -105.2225
11.0508 | -105.2256
11.0545 | -105.2302
11.0581 | -105.2346
11.0618 | -105.2348
11.0655 | -105.2297
11.0692 | -105.2218
11.0729 | -105.2089
11.0766 | -105.1961
11.0803 | -105.1836
11.084 | -105.1663
11.0876 | -105.1471
11.0913 | -105.1296
11.095 | -105.1071
11.0987 | -105.0826
11.1024 | -105.0586
11.1061 | -105.0361
11.1098 | -105.0151
11.1135 | -104.9936
11.1172 | -104.9724
11.1208 | -104.9499
11.1245 | -104.9332
11.1282 | -104.9242
11.1319 | -104.9157
11.1356 | -104.9064
11.1393 | -104.901
11.143 | -104.8923
11.1467 | -104.882
11.1504 | -104.8732
11.154 | -104.869
11.1577 | -104.8707
11.1614 | -104.8705
11.1651 | -104.8711
11.1688 | -104.8724
11.1725 | -104.8806
11.1762 | -104.8863
11.1799 | -104.8903
11.1835 | -104.895
11.1872 | -104.9028
11.1909 | -104.9146
11.1946 | -104.9274
11.1983 | -104.9371
11.202 | -104.9451
11.2057 | -104.9548
11.2094 | -104.9664
11.2131 | -104.9768
11.2167 | -104.9878
11.2204 | -105.0003
11.2241 | -105.0146
11.2278 | -105.03
11.2315 | -105.0454
11.2352 | -105.0631
11.2389 | -105.0817
11.2426 | -105.0989
11.2463 | -105.1152
11.2499 | -105.1293
11.2536 | -105.1406
11.2573 | -105.152
11.261 | -105.1617
11.2647 | -105.1697
11.2684 | -105.177
11.2721 | -105.1842
11.2758 | -105.1893
11.2795 | -105.1932
11.2831 | -105.1957
11.2868 | -105.1944
11.2905 | -105.1864
11.2942 | -105.1783
11.2979 | -105.1735
11.3016 | -105.1665
11.3053 | -105.1607
11.309 | -105.1562
11.3126 | -105.1495
11.3163 | -105.141
11.32 | -105.1302
11.3237 | -105.1222
11.3274 | -105.1159
11.3311 | -105.1091
11.3348 | -105.1003
11.3385 | -105.0943
11.3422 | -105.0911
11.3458 | -105.0883
11.3495 | -105.0868
11.3532 | -105.0875
11.3569 | -105.0877
11.3606 | -105.0919
11.3643 | -105.0966
11.368 | -105.1028
11.3717 | -105.1075
11.3754 | -105.1105
11.379 | -105.1106
11.3827 | -105.1087
11.3864 | -105.1081
11.3901 | -105.1079
11.3938 | -105.1071
11.3975 | -105.1036
11.4012 | -105.0973
11.4049 | -105.0919
11.4085 | -105.0881
11.4122 | -105.0873
11.4159 | -105.0888
11.4196 | -105.092
11.4233 | -105.0974
11.427 | -105.0997
11.4307 | -105.1016
11.4344 | -105.1048
11.4381 | -105.1091
11.4417 | -105.1148
11.4454 | -105.1226
11.4491 | -105.1354
11.4528 | -105.149
11.4565 | -105.1635
11.4602 | -105.1762
11.4639 | -105.1917
11.4676 | -105.207
11.4713 | -105.2205
11.4749 | -105.23
11.4786 | -105.2441
11.4823 | -105.2623
11.486 | -105.282
11.4897 | -105.297
11.4934 | -105.311
11.4971 | -105.3225
11.5008 | -105.336
11.5044 | -105.3489
11.5081 | -105.3607
11.5118 | -105.3741
11.5155 | -105.3846
11.5192 | -105.3914
11.5229 | -105.3989
11.5266 | -105.4003
11.5303 | -105.4067
11.534 | -105.4081
11.5376 | -105.4095
11.5413 | -105.4106
11.545 | -105.4113
11.5487 | -105.4139
11.5524 | -105.4172
11.5561 | -105.4204
11.5598 | -105.4228
11.5635 | -105.4284
11.5672 | -105.4344
11.5708 | -105.4364
11.5745 | -105.4429
11.5782 | -105.4496
11.5819 | -105.4561
11.5856 | -105.4674
11.5893 | -105.4755
11.593 | -105.4877
11.5967 | -105.5023
11.6004 | -105.5156
11.604 | -105.5283
11.6077 | -105.5418
11.6114 | -105.5559
11.6151 | -105.5628
11.6188 | -105.5684
11.6225 | -105.5754
11.6262 | -105.5816
11.6299 | -105.5805
11.6335 | -105.5765
11.6372 | -105.5705
11.6409 | -105.5636
11.6446 | -105.5523
11.6483 | -105.5389
11.652 | -105.5235
11.6557 | -105.5081
11.6594 | -105.4936
11.6631 | -105.4802
11.6667 | -105.4645
11.6704 | -105.4466
11.6741 | -105.4313
11.6778 | -105.4189
11.6815 | -105.4052
11.6852 | -105.3961
11.6889 | -105.3878
11.6926 | -105.3864
11.6963 | -105.3848
11.6999 | -105.3847
11.7036 | -105.379
11.7073 | -105.3755
11.711 | -105.3696
11.7147 | -105.3643
11.7184 | -105.3573
11.7221 | -105.3479
11.7258 | -105.3379
11.7294 | -105.3275
11.7331 | -105.3145
11.7368 | -105.2998
11.7405 | -105.2861
11.7442 | -105.2662
11.7479 | -105.2449
11.7516 | -105.223
11.7553 | -105.1973
11.759 | -105.176
11.7626 | -105.1549
11.7663 | -105.1314
11.77 | -105.1066
11.7737 | -105.0836
11.7774 | -105.0646
11.7811 | -105.0457
11.7848 | -105.0327
11.7885 | -105.0216
11.7922 | -105.0142
11.7958 | -105.0083
11.7995 | -105.006
11.8032 | -105.0064
11.8069 | -105.0081
11.8106 | -105.0142
11.8143 | -105.0209
11.818 | -105.0294
11.8217 | -105.0406
11.8253 | -105.0525
11.829 | -105.0642
11.8327 | -105.0786
11.8364 | -105.0938
11.8401 | -105.1076
11.8438 | -105.1181
11.8475 | -105.1287
11.8512 | -105.1444
11.8549 | -105.158
11.8585 | -105.1718
11.8622 | -105.1842
11.8659 | -105.1924
11.8696 | -105.2043
11.8733 | -105.2175
11.877 | -105.2296
11.8807 | -105.2408
11.8844 | -105.2478
11.8881 | -105.2539
11.8917 | -105.2617
11.8954 | -105.2654
11.8991 | -105.2693
11.9028 | -105.2757
11.9065 | -105.2794
11.9102 | -105.2811
11.9139 | -105.2838
11.9176 | -105.2861
11.9213 | -105.2895
11.9249 | -105.295
11.9286 | -105.3029
11.9323 | -105.3037
11.936 | -105.304
11.9397 | -105.3085
11.9434 | -105.3126
11.9471 | -105.3157
11.9508 | -105.3197
11.9544 | -105.3235
11.9581 | -105.3261
11.9618 | -105.3302
11.9655 | -105.3327
11.9692 | -105.3366
11.9729 | -105.3398
11.9766 | -105.3421
11.9803 | -105.3455
11.984 | -105.3453
11.9876 | -105.3468
11.9913 | -105.3489
11.995 | -105.3517
11.9987 | -105.3574
12.0024 | -105.3617
12.0061 | -105.3607
12.0098 | -105.3607
12.0135 | -105.3643
12.0172 | -105.3661
12.0208 | -105.3652
12.0245 | -105.3616
12.0282 | -105.3583
12.0319 | -105.3523
12.0356 | -105.3438
12.0393 | -105.341
12.043 | -105.3337
12.0467 | -105.3243
12.0503 | -105.3148
12.054 | -105.3066
12.0577 | -105.3056
12.0614 | -105.3028
12.0651 | -105.2969
12.0688 | -105.2964
12.0725 | -105.2966
12.0762 | -105.2989
12.0799 | -105.307
12.0835 | -105.3128
12.0872 | -105.3155
12.0909 | -105.3207
12.0946 | -105.3249
12.0983 | -105.3254
12.102 | -105.325
12.1057 | -105.3334
12.1094 | -105.3399
12.1131 | -105.3416
12.1167 | -105.3428
12.1204 | -105.3438
12.1241 | -105.3407
12.1278 | -105.3389
12.1315 | -105.3374
12.1352 | -105.3354
12.1389 | -105.3348
12.1426 | -105.3343
12.1462 | -105.3328
12.1499 | -105.3323
12.1536 | -105.3343
12.1573 | -105.3379
12.161 | -105.3404
12.1647 | -105.3431
12.1684 | -105.349
12.1721 | -105.3576
12.1758 | -105.3684
12.1794 | -105.3844
12.1831 | -105.4028
12.1868 | -105.4192
12.1905 | -105.4381
12.1942 | -105.4585
12.1979 | -105.4803
12.2016 | -105.5039
12.2053 | -105.5293
12.209 | -105.5567
12.2126 | -105.5828
12.2163 | -105.6094
12.22 | -105.6345
12.2237 | -105.6582
12.2274 | -105.6767
12.2311 | -105.6926
12.2348 | -105.7081
12.2385 | -105.7116
12.2422 | -105.7245
12.2458 | -105.7357
12.2495 | -105.7419
12.2532 | -105.7436
12.2569 | -105.7464
12.2606 | -105.7516
12.2643 | -105.758
12.268 | -105.7631
12.2717 | -105.7674
12.2753 | -105.7687
12.279 | -105.7652
12.2827 | -105.7582
12.2864 | -105.7518
12.2901 | -105.7487
12.2938 | -105.7467
12.2975 | -105.7395
12.3012 | -105.7337
12.3049 | -105.7285
12.3085 | -105.7222
12.3122 | -105.7121
12.3159 | -105.7016
12.3196 | -105.693
12.3233 | -105.6855
12.327 | -105.675
12.3307 | -105.6623
12.3344 | -105.6507
12.3381 | -105.6393
12.3417 | -105.6254
12.3454 | -105.6084
12.3491 | -105.5918
12.3528 | -105.5752
12.3565 | -105.5581
12.3602 | -105.5408
12.3639 | -105.5243
12.3676 | -105.51
12.3712 | -105.4946
12.3749 | -105.4789
12.3786 | -105.4624
12.3823 | -105.4482
12.386 | -105.4375
12.3897 | -105.4268
12.3934 | -105.4187
12.3971 | -105.4128
12.4008 | -105.4062
12.4044 | -105.3999
12.4081 | -105.3899
12.4118 | -105.3849
12.4155 | -105.3807
12.4192 | -105.379
12.4229 | -105.3776
12.4266 | -105.3765
12.4303 | -105.3779
12.434 | -105.3807
12.4376 | -105.3809
12.4413 | -105.382
12.445 | -105.3846
12.4487 | -105.3869
12.4524 | -105.3899
12.4561 | -105.3954
12.4598 | -105.3986
12.4635 | -105.4016
12.4671 | -105.4058
12.4708 | -105.4072
12.4745 | -105.409
12.4782 | -105.4146
12.4819 | -105.4195
12.4856 | -105.4193
12.4893 | -105.419
12.493 | -105.4232
12.4967 | -105.4263
12.5003 | -105.4273
12.504 | -105.4298
12.5077 | -105.4359
12.5114 | -105.4402
12.5151 | -105.4461
12.5188 | -105.455
12.5225 | -105.4661
12.5262 | -105.4745
12.5299 | -105.4824
12.5335 | -105.4892
12.5372 | -105.4969
12.5409 | -105.5037
12.5446 | -105.5118
12.5483 | -105.5217
12.552 | -105.5332
12.5557 | -105.5434
12.5594 | -105.557
12.5631 | -105.5719
12.5667 | -105.5852
12.5704 | -105.6009
12.5741 | -105.6182
12.5778 | -105.6382
12.5815 | -105.659
12.5852 | -105.6785
12.5889 | -105.7027
12.5926 | -105.7268
12.5962 | -105.7534
12.5999 | -105.779
12.6036 | -105.8032
12.6073 | -105.827
12.611 | -105.8499
12.6147 | -105.8701
12.6184 | -105.8918
12.6221 | -105.9096
12.6258 | -105.9255
12.6294 | -105.9432
12.6331 | -105.9607
12.6368 | -105.9791
12.6405 | -105.9956
12.6442 | -106.006
12.6479 | -106.0153
12.6516 | -106.0207
12.6553 | -106.0219
12.659 | -106.021
12.6626 | -106.0154
12.6663 | -106.0126
12.67 | -106.0075
12.6737 | -105.9968
12.6774 | -105.9825
12.6811 | -105.9689
12.6848 | -105.9524
12.6885 | -105.9348
12.6921 | -105.9168
12.6958 | -105.8956
12.6995 | -105.8732
12.7032 | -105.8504
12.7069 | -105.8329
12.7106 | -105.8127
12.7143 | -105.7953
12.718 | -105.7779
12.7217 | -105.7598
12.7253 | -105.7434
12.729 | -105.7292
12.7327 | -105.7136
12.7364 | -105.7005
12.7401 | -105.6898
12.7438 | -105.6786
12.7475 | -105.6693
12.7512 | -105.6652
12.7549 | -105.6635
12.7585 | -105.6643
12.7622 | -105.6674
12.7659 | -105.6743
12.7696 | -105.6753
12.7733 | -105.6775
12.777 | -105.6751
12.7807 | -105.6709
12.7844 | -105.6685
12.788 | -105.6668
12.7917 | -105.6652
12.7954 | -105.6619
12.7991 | -105.6566
12.8028 | -105.6503
12.8065 | -105.6451
12.8102 | -105.635
12.8139 | -105.6207
12.8176 | -105.607
12.8212 | -105.5923
12.8249 | -105.5798
12.8286 | -105.5688
12.8323 | -105.5547
12.836 | -105.545
12.8397 | -105.5366
12.8434 | -105.5262
12.8471 | -105.5193
12.8508 | -105.513
12.8544 | -105.5087
12.8581 | -105.5053
12.8618 | -105.5051
12.8655 | -105.5048
12.8692 | -105.5017
12.8729 | -105.4993
12.8766 | -105.4989
12.8803 | -105.4992
12.884 | -105.4979
12.8876 | -105.4961
12.8913 | -105.4913
12.895 | -105.4882
12.8987 | -105.4849
12.9024 | -105.482
12.9061 | -105.4809
12.9098 | -105.4763
12.9135 | -105.4697
12.9171 | -105.4665
12.9208 | -105.4617
12.9245 | -105.4516
12.9282 | -105.4464
12.9319 | -105.444
12.9356 | -105.439
12.9393 | -105.4347
12.943 | -105.4329
12.9467 | -105.4301
12.9503 | -105.4244
12.954 | -105.4243
12.9577 | -105.4236
12.9614 | -105.4237
12.9651 | -105.4263
12.9688 | -105.4275
12.9725 | -105.4317
12.9762 | -105.4381
12.9799 | -105.4449
12.9835 | -105.4509
12.9872 | -105.4566
12.9909 | -105.4613
12.9946 | -105.4658
12.9983 | -105.4681
13.002 | -105.4686
13.0057 | -105.4697
13.0094 | -105.4724
13.013 | -105.4728
13.0167 | -105.4685
13.0204 | -105.4651
13.0241 | -105.4594
13.0278 | -105.4528
13.0315 | -105.4471
13.0352 | -105.4372
13.0389 | -105.427
13.0426 | -105.4178
13.0462 | -105.4095
13.0499 | -105.3987
13.0536 | -105.3874
13.0573 | -105.382
13.061 | -105.3758
13.0647 | -105.3689
13.0684 | -105.3616
13.0721 | -105.3599
13.0758 | -105.3614
13.0794 | -105.3644
13.0831 | -105.3678
13.0868 | -105.372
13.0905 | -105.3778
13.0942 | -105.3857
13.0979 | -105.3981
13.1016 | -105.4099
13.1053 | -105.4222
13.1089 | -105.4396
13.1126 | -105.4583
13.1163 | -105.4778
13.12 | -105.495
13.1237 | -105.5122
13.1274 | -105.5285
13.1311 | -105.5439
13.1348 | -105.5634
13.1385 | -105.587
13.1421 | -105.6064
13.1458 | -105.6275
13.1495 | -105.6467
13.1532 | -105.6633
13.1569 | -105.6817
13.1606 | -105.6958
13.1643 | -105.7031
13.168 | -105.7102
13.1717 | -105.7178
13.1753 | -105.7252
13.179 | -105.7282
13.1827 | -105.7259
13.1864 | -105.7212
13.1901 | -105.7178
13.1938 | -105.7102
13.1975 | -105.6997
13.2012 | -105.6893
13.2049 | -105.6819
13.2085 | -105.668
13.2122 | -105.6524
13.2159 | -105.6385
13.2196 | -105.6231
13.2233 | -105.6045
13.227 | -105.5843
13.2307 | -105.567
13.2344 | -105.5525
13.238 | -105.5397
13.2417 | -105.5273
13.2454 | -105.5115
13.2491 | -105.4925
13.2528 | -105.4768
13.2565 | -105.4615
13.2602 | -105.4445
13.2639 | -105.4292
13.2676 | -105.415
13.2712 | -105.4016
13.2749 | -105.3909
13.2786 | -105.3811
13.2823 | -105.3707
13.286 | -105.3622
13.2897 | -105.3571
13.2934 | -105.3519
13.2971 | -105.3441
13.3008 | -105.3461
13.3044 | -105.3435
13.3081 | -105.3372
13.3118 | -105.3336
13.3155 | -105.3344
13.3192 | -105.3371
13.3229 | -105.3391
13.3266 | -105.3403
13.3303 | -105.3408
13.3339 | -105.3426
13.3376 | -105.3439
13.3413 | -105.347
13.345 | -105.3509
13.3487 | -105.3568
13.3524 | -105.3617
13.3561 | -105.3694
13.3598 | -105.3759
13.3635 | -105.3787
13.3671 | -105.3797
13.3708 | -105.3809
13.3745 | -105.3816
13.3782 | -105.3816
13.3819 | -105.3793
13.3856 | -105.3711
13.3893 | -105.3617
13.393 | -105.3511
13.3967 | -105.3379
13.4003 | -105.3294
13.404 | -105.3187
13.4077 | -105.3045
13.4114 | -105.2932
13.4151 | -105.2846
13.4188 | -105.2757
13.4225 | -105.2619
13.4262 | -105.2483
13.4298 | -105.2368
13.4335 | -105.2281
13.4372 | -105.2238
13.4409 | -105.2231
13.4446 | -105.2206
13.4483 | -105.2189
13.452 | -105.2172
13.4557 | -105.2164
13.4594 | -105.2172
13.463 | -105.2186
13.4667 | -105.2193
13.4704 | -105.2235
13.4741 | -105.2309
13.4778 | -105.2402
13.4815 | -105.2489
13.4852 | -105.2612
13.4889 | -105.2784
13.4926 | -105.2948
13.4962 | -105.3118
13.4999 | -105.3319
13.5036 | -105.3568
13.5073 | -105.3808
13.511 | -105.4054
13.5147 | -105.4333
13.5184 | -105.4617
13.5221 | -105.4894
13.5258 | -105.5178
13.5294 | -105.5462
13.5331 | -105.5744
13.5368 | -105.6015
13.5405 | -105.6291
13.5442 | -105.6523
13.5479 | -105.673
13.5516 | -105.6951
13.5553 | -105.7184
13.5589 | -105.7407
13.5626 | -105.7631
13.5663 | -105.7817
13.57 | -105.7971
13.5737 | -105.8102
13.5774 | -105.82
13.5811 | -105.8307
13.5848 | -105.8391
13.5885 | -105.8468
13.5921 | -105.8544
13.5958 | -105.8551
13.5995 | -105.8541
13.6032 | -105.8508
13.6069 | -105.8468
13.6106 | -105.8429
13.6143 | -105.8365
13.618 | -105.8295
13.6217 | -105.8224
13.6253 | -105.8158
13.629 | -105.805
13.6327 | -105.7915
13.6364 | -105.7801
13.6401 | -105.7722
13.6438 | -105.7647
13.6475 | -105.7547
13.6512 | -105.745
13.6548 | -105.7372
13.6585 | -105.7306
13.6622 | -105.725
13.6659 | -105.7182
13.6696 | -105.7109
13.6733 | -105.7051
13.677 | -105.7035
13.6807 | -105.7027
13.6844 | -105.7023
13.688 | -105.7057
13.6917 | -105.7099
13.6954 | -105.7142
13.6991 | -105.7189
13.7028 | -105.7248
13.7065 | -105.7283
13.7102 | -105.7308
13.7139 | -105.7369
13.7176 | -105.7446
13.7212 | -105.7531
13.7249 | -105.7609
13.7286 | -105.7671
13.7323 | -105.7733
13.736 | -105.7771
13.7397 | -105.7815
13.7434 | -105.79
13.7471 | -105.7938
13.7507 | -105.7949
13.7544 | -105.7977
13.7581 | -105.7968
13.7618 | -105.7975
13.7655 | -105.8015
13.7692 | -105.8043
13.7729 | -105.8042
13.7766 | -105.806
13.7803 | -105.8084
13.7839 | -105.8101
13.7876 | -105.8138
13.7913 | -105.8206
13.795 | -105.825
13.7987 | -105.8308
13.8024 | -105.835
13.8061 | -105.8398
13.8098 | -105.8425
13.8135 | -105.8464
13.8171 | -105.851
13.8208 | -105.8545
13.8245 | -105.8622
13.8282 | -105.868
13.8319 | -105.8723
13.8356 | -105.8795
13.8393 | -105.8843
13.843 | -105.8905
13.8467 | -105.9009
13.8503 | -105.9127
13.854 | -105.9244
13.8577 | -105.9377
13.8614 | -105.9495
13.8651 | -105.961
13.8688 | -105.9749
13.8725 | -105.9864
13.8762 | -105.9977
13.8798 | -106.0087
13.8835 | -106.019
13.8872 | -106.0291
13.8909 | -106.0374
13.8946 | -106.0457
13.8983 | -106.053
13.902 | -106.0617
13.9057 | -106.0716
13.9094 | -106.0771
13.913 | -106.0772
13.9167 | -106.0785
13.9204 | -106.0778
13.9241 | -106.0748
13.9278 | -106.0722
13.9315 | -106.0687
13.9352 | -106.0627
13.9389 | -106.0551
13.9426 | -106.0453
13.9462 | -106.0338
13.9499 | -106.0208
13.9536 | -106.0062
13.9573 | -105.9903
13.961 | -105.9714
13.9647 | -105.9503
13.9684 | -105.9309
13.9721 | -105.9105
13.9757 | -105.8885
13.9794 | -105.8669
13.9831 | -105.8458
13.9868 | -105.8222
13.9905 | -105.8001
13.9942 | -105.7839
13.9979 | -105.767
14.0016 | -105.7514
14.0053 | -105.7367
14.0089 | -105.7239
14.0126 | -105.7149
14.0163 | -105.7086
14.02 | -105.7049
14.0237 | -105.703
14.0274 | -105.7051
14.0311 | -105.7096
14.0348 | -105.7138
14.0385 | -105.718
14.0421 | -105.7254
14.0458 | -105.735
14.0495 | -105.7438
14.0532 | -105.754
14.0569 | -105.7672
14.0606 | -105.7816
14.0643 | -105.7977
14.068 | -105.8138
14.0716 | -105.8286
14.0753 | -105.843
14.079 | -105.8584
14.0827 | -105.8703
14.0864 | -105.8774
14.0901 | -105.8851
14.0938 | -105.8956
14.0975 | -105.9037
14.1012 | -105.9078
14.1048 | -105.9081
14.1085 | -105.9038
14.1122 | -105.9009
14.1159 | -105.8989
14.1196 | -105.8927
14.1233 | -105.8866
14.127 | -105.8837
14.1307 | -105.8796
14.1344 | -105.8751
14.138 | -105.8726
14.1417 | -105.8693
14.1454 | -105.8627
14.1491 | -105.8582
14.1528 | -105.8537
14.1565 | -105.8495
14.1602 | -105.8469
14.1639 | -105.8463
14.1676 | -105.8457
14.1712 | -105.8429
14.1749 | -105.8444
14.1786 | -105.8462
14.1823 | -105.848
14.186 | -105.8497
14.1897 | -105.8521
14.1934 | -105.8566
14.1971 | -105.8593
14.2007 | -105.8621
14.2044 | -105.8634
14.2081 | -105.8681
14.2118 | -105.8746
14.2155 | -105.8805
14.2192 | -105.8865
14.2229 | -105.8901
14.2266 | -105.8927
14.2303 | -105.8984
14.2339 | -105.9048
14.2376 | -105.9124
14.2413 | -105.9218
14.245 | -105.9279
14.2487 | -105.935
14.2524 | -105.9428
14.2561 | -105.9469
14.2598 | -105.9516
14.2635 | -105.9594
14.2671 | -105.969
14.2708 | -105.9789
14.2745 | -105.9873
14.2782 | -105.996
14.2819 | -106.0038
14.2856 | -106.0136
14.2893 | -106.0267
14.293 | -106.0377
14.2966 | -106.0457
14.3003 | -106.0549
14.304 | -106.0666
14.3077 | -106.0804
14.3114 | -106.0933
14.3151 | -106.1038
14.3188 | -106.1163
14.3225 | -106.1274
14.3262 | -106.1371
14.3298 | -106.1428
14.3335 | -106.1465
14.3372 | -106.1502
14.3409 | -106.1544
14.3446 | -106.1601
14.3483 | -106.1682
14.352 | -106.1751
14.3557 | -106.1809
14.3594 | -106.185
14.363 | -106.1863
14.3667 | -106.1813
14.3704 | -106.1807
14.3741 | -106.1796
14.3778 | -106.1781
14.3815 | -106.174
14.3852 | -106.172
14.3889 | -106.1708
14.3925 | -106.1693
14.3962 | -106.1667
14.3999 | -106.1631
14.4036 | -106.1604
14.4073 | -106.1558
14.411 | -106.1513
14.4147 | -106.1475
14.4184 | -106.1441
14.4221 | -106.1383
14.4257 | -106.1307
14.4294 | -106.1237
14.4331 | -106.1208
14.4368 | -106.1162
14.4405 | -106.112
14.4442 | -106.1058
14.4479 | -106.0995
14.4516 | -106.0971
14.4553 | -106.0919
14.4589 | -106.0896
14.4626 | -106.0894
14.4663 | -106.0849
14.47 | -106.0803
14.4737 | -106.0777
14.4774 | -106.0731
14.4811 | -106.0702
14.4848 | -106.0684
14.4885 | -106.0677
14.4921 | -106.0676
14.4958 | -106.0688
14.4995 | -106.0692
14.5032 | -106.0686
14.5069 | -106.0686
14.5106 | -106.071
14.5143 | -106.0732
14.518 | -106.0773
14.5216 | -106.0809
14.5253 | -106.0823
14.529 | -106.0846
14.5327 | -106.0883
14.5364 | -106.0929
14.5401 | -106.0972
14.5438 | -106.1021
14.5475 | -106.1036
14.5512 | -106.1054
14.5548 | -106.1038
14.5585 | -106.1042
14.5622 | -106.1061
14.5659 | -106.1081
14.5696 | -106.1117
14.5733 | -106.115
14.577 | -106.1173
14.5807 | -106.1181
14.5844 | -106.1184
14.588 | -106.1192
14.5917 | -106.1201
14.5954 | -106.1179
14.5991 | -106.116
14.6028 | -106.1188
14.6065 | -106.1216
14.6102 | -106.1246
14.6139 | -106.1276
14.6175 | -106.1356
14.6212 | -106.144
14.6249 | -106.1492
14.6286 | -106.1574
14.6323 | -106.1663
14.636 | -106.1749
14.6397 | -106.1854
14.6434 | -106.1929
14.6471 | -106.2006
14.6507 | -106.207
14.6544 | -106.2137
14.6581 | -106.2209
14.6618 | -106.2268
14.6655 | -106.2331
14.6692 | -106.2372
14.6729 | -106.2405
14.6766 | -106.2454
14.6803 | -106.2483
14.6839 | -106.251
14.6876 | -106.2586
14.6913 | -106.263
14.695 | -106.2649
14.6987 | -106.2654
14.7024 | -106.2683
14.7061 | -106.2704
14.7098 | -106.2732
14.7134 | -106.2791
14.7171 | -106.284
14.7208 | -106.2845
14.7245 | -106.2843
14.7282 | -106.2855
14.7319 | -106.2892
14.7356 | -106.2936
14.7393 | -106.297
14.743 | -106.3005
14.7466 | -106.3046
14.7503 | -106.3092
14.754 | -106.3139
14.7577 | -106.3201
14.7614 | -106.3244
14.7651 | -106.3314
14.7688 | -106.3362
14.7725 | -106.3411
14.7762 | -106.3464
14.7798 | -106.3528
14.7835 | -106.3605
14.7872 | -106.3668
14.7909 | -106.374
14.7946 | -106.3807
14.7983 | -106.383
14.802 | -106.3867
14.8057 | -106.3923
14.8094 | -106.3968
14.813 | -106.3976
14.8167 | -106.399
14.8204 | -106.3949
14.8241 | -106.3928
14.8278 | -106.3874
14.8315 | -106.381
14.8352 | -106.3742
14.8389 | -106.3655
14.8425 | -106.3564
14.8462 | -106.3456
14.8499 | -106.3336
14.8536 | -106.3195
14.8573 | -106.3066
14.861 | -106.2957
14.8647 | -106.2852
14.8684 | -106.272
14.8721 | -106.2582
14.8757 | -106.2451
14.8794 | -106.2328
14.8831 | -106.2192
14.8868 | -106.2066
14.8905 | -106.1977
14.8942 | -106.1883
14.8979 | -106.1876
14.9016 | -106.1783
14.9053 | -106.1736
14.9089 | -106.1714
14.9126 | -106.1681
14.9163 | -106.1645
14.92 | -106.1615
14.9237 | -106.1607
14.9274 | -106.1597
14.9311 | -106.1613
14.9348 | -106.1665
14.9384 | -106.1738
14.9421 | -106.1813
14.9458 | -106.1832
14.9495 | -106.1854
14.9532 | -106.1915
14.9569 | -106.1969
14.9606 | -106.2038
14.9643 | -106.2122
14.968 | -106.2148
14.9716 | -106.2175
14.9753 | -106.2195
14.979 | -106.2186
14.9827 | -106.2169
14.9864 | -106.2151
14.9901 | -106.2117
14.9938 | -106.205
14.9975 | -106.1977
15.0012 | -106.1912
15.0048 | -106.1839
15.0085 | -106.1747
15.0122 | -106.1667
15.0159 | -106.1567
15.0196 | -106.1451
15.0233 | -106.1277
15.027 | -106.1074
15.0307 | -106.0901
15.0343 | -106.0749
15.038 | -106.0559
15.0417 | -106.0387
15.0454 | -106.0232
15.0491 | -106.0088
15.0528 | -105.9953
15.0565 | -105.9829
15.0602 | -105.9733
15.0639 | -105.9652
15.0675 | -105.9582
15.0712 | -105.95
15.0749 | -105.9438
15.0786 | -105.9378
15.0823 | -105.937
15.086 | -105.9386
15.0897 | -105.9446
15.0934 | -105.9507
15.0971 | -105.9604
15.1007 | -105.9682
15.1044 | -105.9714
15.1081 | -105.9772
15.1118 | -105.9835
15.1155 | -105.9911
15.1192 | -105.9983
15.1229 | -106.0039
15.1266 | -106.0113
15.1303 | -106.0205
15.1339 | -106.0288
15.1376 | -106.0362
15.1413 | -106.0414
15.145 | -106.045
15.1487 | -106.0463
15.1524 | -106.048
15.1561 | -106.0479
15.1598 | -106.0459
15.1634 | -106.0488
15.1671 | -106.0486
15.1708 | -106.0453
15.1745 | -106.0411
15.1782 | -106.0354
15.1819 | -106.0288
15.1856 | -106.0249
15.1893 | -106.0215
15.193 | -106.0167
15.1966 | -106.0126
15.2003 | -106.0089
15.204 | -106.005
15.2077 | -106.0028
15.2114 | -106.0025
15.2151 | -106.0034
15.2188 | -106.0047
15.2225 | -106.004
15.2262 | -106.0057
15.2298 | -106.0084
15.2335 | -106.0099
15.2372 | -106.0147
15.2409 | -106.0254
15.2446 | -106.0384
15.2483 | -106.0476
15.252 | -106.0569
15.2557 | -106.0668
15.2593 | -106.0752
15.263 | -106.0792
15.2667 | -106.0839
15.2704 | -106.0881
15.2741 | -106.0937
15.2778 | -106.1
15.2815 | -106.102
15.2852 | -106.1021
15.2889 | -106.1023
15.2925 | -106.1002
15.2962 | -106.0985
15.2999 | -106.0918
15.3036 | -106.0841
15.3073 | -106.0732
15.311 | -106.062
15.3147 | -106.0509
15.3184 | -106.0394
15.3221 | -106.0283
15.3257 | -106.0159
15.3294 | -106.0033
15.3331 | -105.992
15.3368 | -105.9824
15.3405 | -105.971
15.3442 | -105.9581
15.3479 | -105.9432
15.3516 | -105.9311
15.3552 | -105.9194
15.3589 | -105.9094
15.3626 | -105.9004
15.3663 | -105.8897
15.37 | -105.8779
15.3737 | -105.8679
15.3774 | -105.8589
15.3811 | -105.8502
15.3848 | -105.8421
15.3884 | -105.8343
15.3921 | -105.8276
15.3958 | -105.821
15.3995 | -105.8154
15.4032 | -105.8109
15.4069 | -105.8056
15.4106 | -105.802
15.4143 | -105.7965
15.418 | -105.79
15.4216 | -105.786
15.4253 | -105.7814
15.429 | -105.7866
15.4327 | -105.7855
15.4364 | -105.7842
15.4401 | -105.783
15.4438 | -105.7851
15.4475 | -105.7857
15.4512 | -105.7889
15.4548 | -105.795
15.4585 | -105.7984
15.4622 | -105.8019
15.4659 | -105.8057
15.4696 | -105.8104
15.4733 | -105.8169
15.477 | -105.8254
15.4807 | -105.8358
15.4843 | -105.8471
15.488 | -105.8546
15.4917 | -105.8623
15.4954 | -105.8699
15.4991 | -105.8787
15.5028 | -105.889
15.5065 | -105.9005
15.5102 | -105.9105
15.5139 | -105.9203
15.5175 | -105.9313
15.5212 | -105.9419
15.5249 | -105.9517
15.5286 | -105.9633
15.5323 | -105.9737
15.536 | -105.9825
15.5397 | -105.9938
15.5434 | -106.0059
15.5471 | -106.0171
15.5507 | -106.0282
15.5544 | -106.0376
15.5581 | -106.0443
15.5618 | -106.0513
15.5655 | -106.0556
15.5692 | -106.0657
15.5729 | -106.0747
15.5766 | -106.0862
15.5802 | -106.0972
15.5839 | -106.1071
15.5876 | -106.1191
15.5913 | -106.1319
15.595 | -106.1417
15.5987 | -106.1513
15.6024 | -106.16
15.6061 | -106.1674
15.6098 | -106.1744
15.6134 | -106.1835
15.6171 | -106.1917
15.6208 | -106.1993
15.6245 | -106.2051
15.6282 | -106.2106
15.6319 | -106.2159
15.6356 | -106.2239
15.6393 | -106.2305
15.643 | -106.2315
15.6466 | -106.2308
15.6503 | -106.2276
15.654 | -106.2245
15.6577 | -106.2228
15.6614 | -106.2216
15.6651 | -106.2203
15.6688 | -106.2195
15.6725 | -106.2194
15.6761 | -106.2151
15.6798 | -106.2082
15.6835 | -106.2046
15.6872 | -106.2015
15.6909 | -106.1976
15.6946 | -106.1924
15.6983 | -106.1857
15.702 | -106.1787
15.7057 | -106.1713
15.7093 | -106.1665
15.713 | -106.1638
15.7167 | -106.1577
15.7204 | -106.15
15.7241 | -106.1433
15.7278 | -106.1369
15.7315 | -106.1276
15.7352 | -106.1191
15.7389 | -106.1124
15.7425 | -106.106
15.7462 | -106.1042
15.7499 | -106.1012
15.7536 | -106.0967
15.7573 | -106.0955
15.761 | -106.096
15.7647 | -106.0956
15.7684 | -106.0939
15.7721 | -106.0937
15.7757 | -106.0942
15.7794 | -106.0961
15.7831 | -106.1026
15.7868 | -106.1106
15.7905 | -106.1203
15.7942 | -106.1254
15.7979 | -106.1312
15.8016 | -106.1351
15.8052 | -106.1399
15.8089 | -106.1483
15.8126 | -106.158
15.8163 | -106.167
15.82 | -106.1769
15.8237 | -106.1856
15.8274 | -106.1939
15.8311 | -106.2049
15.8348 | -106.2146
15.8384 | -106.223
15.8421 | -106.23
15.8458 | -106.2369
15.8495 | -106.2443
15.8532 | -106.2488
15.8569 | -106.2508
15.8606 | -106.2538
15.8643 | -106.2585
15.868 | -106.2644
15.8716 | -106.2704
15.8753 | -106.2739
15.879 | -106.2775
15.8827 | -106.2779
15.8864 | -106.2792
15.8901 | -106.2793
15.8938 | -106.2774
15.8975 | -106.2744
15.9011 | -106.27
15.9048 | -106.2649
15.9085 | -106.2581
15.9122 | -106.2518
15.9159 | -106.245
15.9196 | -106.2374
15.9233 | -106.2279
15.927 | -106.2169
15.9307 | -106.2069
15.9343 | -106.1958
15.938 | -106.1853
15.9417 | -106.1785
15.9454 | -106.172
15.9491 | -106.163
15.9528 | -106.1572
15.9565 | -106.1551
15.9602 | -106.1512
15.9639 | -106.1528
15.9675 | -106.1481
15.9712 | -106.1464
15.9749 | -106.1433
15.9786 | -106.1425
15.9823 | -106.1453
15.986 | -106.1485
15.9897 | -106.1521
15.9934 | -106.1576
15.997 | -106.166
16.0007 | -106.1726
16.0044 | -106.1802
16.0081 | -106.1882
16.0118 | -106.1937
16.0155 | -106.1982
16.0192 | -106.2016
16.0229 | -106.205
16.0266 | -106.2075
16.0302 | -106.2108
16.0339 | -106.2131
16.0376 | -106.2134
16.0413 | -106.2111
16.045 | -106.2071
16.0487 | -106.2057
16.0524 | -106.2052
16.0561 | -106.2023
16.0598 | -106.1973
16.0634 | -106.193
16.0671 | -106.1878
16.0708 | -106.1848
16.0745 | -106.1817
16.0782 | -106.1765
16.0819 | -106.1733
16.0856 | -106.1713
16.0893 | -106.1656
16.0929 | -106.161
16.0966 | -106.1562
16.1003 | -106.1517
16.104 | -106.149
16.1077 | -106.1475
16.1114 | -106.1481
16.1151 | -106.1486
16.1188 | -106.1512
16.1225 | -106.1528
16.1261 | -106.1573
16.1298 | -106.1632
16.1335 | -106.1692
16.1372 | -106.1752
16.1409 | -106.1825
16.1446 | -106.1916
16.1483 | -106.2017
16.152 | -106.2126
16.1557 | -106.2246
16.1593 | -106.2365
16.163 | -106.2478
16.1667 | -106.2591
16.1704 | -106.2685
16.1741 | -106.2761
16.1778 | -106.287
16.1815 | -106.2962
16.1852 | -106.3054
16.1889 | -106.3172
16.1925 | -106.3261
16.1962 | -106.3358
16.1999 | -106.3451
16.2036 | -106.3507
16.2073 | -106.3568
16.211 | -106.3628
16.2147 | -106.3627
16.2184 | -106.3618
16.222 | -106.3622
16.2257 | -106.3611
16.2294 | -106.3564
16.2331 | -106.3507
16.2368 | -106.344
16.2405 | -106.3339
16.2442 | -106.3263
16.2479 | -106.3183
16.2516 | -106.3121
16.2552 | -106.3034
16.2589 | -106.2923
16.2626 | -106.28
16.2663 | -106.2671
16.27 | -106.2552
16.2737 | -106.244
16.2774 | -106.2347
16.2811 | -106.2238
16.2848 | -106.2127
16.2884 | -106.2027
16.2921 | -106.1931
16.2958 | -106.1853
16.2995 | -106.1798
16.3032 | -106.1745
16.3069 | -106.1699
16.3106 | -106.1685
16.3143 | -106.1683
16.3179 | -106.1704
16.3216 | -106.1758
16.3253 | -106.1811
16.329 | -106.1874
16.3327 | -106.1945
16.3364 | -106.2014
16.3401 | -106.2095
16.3438 | -106.2198
16.3475 | -106.2303
16.3511 | -106.246
16.3548 | -106.2645
16.3585 | -106.2803
16.3622 | -106.2959
16.3659 | -106.3139
16.3696 | -106.3335
16.3733 | -106.3511
16.377 | -106.3686
16.3807 | -106.3883
16.3843 | -106.4084
16.388 | -106.4269
16.3917 | -106.446
16.3954 | -106.4658
16.3991 | -106.4835
16.4028 | -106.5008
16.4065 | -106.5169
16.4102 | -106.5283
16.4139 | -106.5381
16.4175 | -106.5458
16.4212 | -106.5539
16.4249 | -106.5615
16.4286 | -106.5655
16.4323 | -106.569
16.436 | -106.5684
16.4397 | -106.5661
16.4434 | -106.5632
16.447 | -106.5598
16.4507 | -106.5536
16.4544 | -106.5445
16.4581 | -106.5364
16.4618 | -106.5286
16.4655 | -106.5189
16.4692 | -106.5072
16.4729 | -106.4974
16.4766 | -106.4834
16.4802 | -106.4697
16.4839 | -106.4513
16.4876 | -106.436
16.4913 | -106.424
16.495 | -106.4173
16.4987 | -106.4066
16.5024 | -106.3931
16.5061 | -106.3806
16.5098 | -106.3692
16.5134 | -106.3581
16.5171 | -106.3452
16.5208 | -106.3357
16.5245 | -106.3278
16.5282 | -106.3202
16.5319 | -106.3122
16.5356 | -106.3018
16.5393 | -106.2929
16.5429 | -106.286
16.5466 | -106.2822
16.5503 | -106.2772
16.554 | -106.2705
16.5577 | -106.2641
16.5614 | -106.2602
16.5651 | -106.2586
16.5688 | -106.2562
16.5725 | -106.2515
16.5761 | -106.2472
16.5798 | -106.2448
16.5835 | -106.2422
16.5872 | -106.2381
16.5909 | -106.2358
16.5946 | -106.2374
16.5983 | -106.243
16.602 | -106.2476
16.6057 | -106.2539
16.6093 | -106.2623
16.613 | -106.2717
16.6167 | -106.2809
16.6204 | -106.2909
16.6241 | -106.3046
16.6278 | -106.318
16.6315 | -106.334
16.6352 | -106.3526
16.6388 | -106.3713
16.6425 | -106.3918
16.6462 | -106.4168
16.6499 | -106.4436
16.6536 | -106.4714
16.6573 | -106.4969
16.661 | -106.5228
16.6647 | -106.5498
16.6684 | -106.5765
16.672 | -106.6054
16.6757 | -106.6341
16.6794 | -106.6647
16.6831 | -106.6956
16.6868 | -106.7242
16.6905 | -106.7504
16.6942 | -106.7729
16.6979 | -106.7938
16.7016 | -106.8149
16.7052 | -106.8337
16.7089 | -106.8483
16.7126 | -106.8634
16.7163 | -106.8777
16.72 | -106.8929
16.7237 | -106.9002
16.7274 | -106.9051
16.7311 | -106.9055
16.7347 | -106.9062
16.7384 | -106.9058
16.7421 | -106.9037
16.7458 | -106.8971
16.7495 | -106.8881
16.7532 | -106.877
16.7569 | -106.866
16.7606 | -106.8492
16.7643 | -106.831
16.7679 | -106.8103
16.7716 | -106.788
16.7753 | -106.7677
16.779 | -106.747
16.7827 | -106.7244
16.7864 | -106.6987
16.7901 | -106.6724
16.7938 | -106.6495
16.7975 | -106.6268
16.8011 | -106.6044
16.8048 | -106.5821
16.8085 | -106.5635
16.8122 | -106.5485
16.8159 | -106.5311
16.8196 | -106.5152
16.8233 | -106.5012
16.827 | -106.4868
16.8307 | -106.4721
16.8343 | -106.4582
16.838 | -106.4458
16.8417 | -106.4375
16.8454 | -106.4307
16.8491 | -106.4253
16.8528 | -106.4216
16.8565 | -106.421
16.8602 | -106.422
16.8638 | -106.4211
16.8675 | -106.421
16.8712 | -106.4224
16.8749 | -106.4258
16.8786 | -106.4299
16.8823 | -106.4357
16.886 | -106.4397
16.8897 | -106.4442
16.8934 | -106.4479
16.897 | -106.4529
16.9007 | -106.4606
16.9044 | -106.4655
16.9081 | -106.4699
16.9118 | -106.476
16.9155 | -106.4831
16.9192 | -106.4883
16.9229 | -106.4943
16.9266 | -106.5022
16.9302 | -106.5117
16.9339 | -106.5233
16.9376 | -106.5337
16.9413 | -106.5451
16.945 | -106.5568
16.9487 | -106.569
16.9524 | -106.5792
16.9561 | -106.5864
16.9597 | -106.5936
16.9634 | -106.6034
16.9671 | -106.6102
16.9708 | -106.6177
16.9745 | -106.6247
16.9782 | -106.6332
16.9819 | -106.6418
16.9856 | -106.649
16.9893 | -106.6555
16.9929 | -106.6641
16.9966 | -106.6701
17.0003 | -106.6745
17.004 | -106.6794
17.0077 | -106.6836
17.0114 | -106.6862
17.0151 | -106.6894
17.0188 | -106.6899
17.0225 | -106.6875
17.0261 | -106.6857
17.0298 | -106.685
17.0335 | -106.6855
17.0372 | -106.6833
17.0409 | -106.6782
17.0446 | -106.6738
17.0483 | -106.6711
17.052 | -106.6667
17.0556 | -106.6607
17.0593 | -106.6556
17.063 | -106.6487
17.0667 | -106.6423
17.0704 | -106.6357
17.0741 | -106.6277
17.0778 | -106.6221
17.0815 | -106.613
17.0852 | -106.6009
17.0888 | -106.5911
17.0925 | -106.5824
17.0962 | -106.576
17.0999 | -106.5696
17.1036 | -106.5631
17.1073 | -106.5575
17.111 | -106.5533
17.1147 | -106.5493
17.1184 | -106.5453
17.122 | -106.5436
17.1257 | -106.54
17.1294 | -106.5373
17.1331 | -106.5387
17.1368 | -106.542
17.1405 | -106.5444
17.1442 | -106.5468
17.1479 | -106.5559
17.1516 | -106.565
17.1552 | -106.5734
17.1589 | -106.5805
17.1626 | -106.5889
17.1663 | -106.6
17.17 | -106.6126
17.1737 | -106.624
17.1774 | -106.6357
17.1811 | -106.6483
17.1847 | -106.661
17.1884 | -106.6744
17.1921 | -106.6863
17.1958 | -106.6995
17.1995 | -106.7124
17.2032 | -106.7239
17.2069 | -106.7356
17.2106 | -106.7454
17.2143 | -106.7535
17.2179 | -106.7593
17.2216 | -106.7632
17.2253 | -106.7663
17.229 | -106.7667
17.2327 | -106.7639
17.2364 | -106.7594
17.2401 | -106.7555
17.2438 | -106.7528
17.2475 | -106.7471
17.2511 | -106.7406
17.2548 | -106.732
17.2585 | -106.7197
17.2622 | -106.7076
17.2659 | -106.6961
17.2696 | -106.6814
17.2733 | -106.6636
17.277 | -106.6475
17.2806 | -106.6321
17.2843 | -106.6188
17.288 | -106.6066
17.2917 | -106.5922
17.2954 | -106.5793
17.2991 | -106.568
17.3028 | -106.5571
17.3065 | -106.5456
17.3102 | -106.5344
17.3138 | -106.5246
17.3175 | -106.5176
17.3212 | -106.5147
17.3249 | -106.5111
17.3286 | -106.5089
17.3323 | -106.509
17.336 | -106.5076
17.3397 | -106.5063
17.3434 | -106.5076
17.347 | -106.5104
17.3507 | -106.5162
17.3544 | -106.5228
17.3581 | -106.5291
17.3618 | -106.5372
17.3655 | -106.5435
17.3692 | -106.5499
17.3729 | -106.5606
17.3766 | -106.5702
17.3802 | -106.5765
17.3839 | -106.5842
17.3876 | -106.5944
17.3913 | -106.6037
17.395 | -106.6082
17.3987 | -106.614
17.4024 | -106.618
17.4061 | -106.6242
17.4097 | -106.6291
17.4134 | -106.6296
17.4171 | -106.6304
17.4208 | -106.6317
17.4245 | -106.6315
17.4282 | -106.6337
17.4319 | -106.6325
17.4356 | -106.6284
17.4393 | -106.6267
17.4429 | -106.6231
17.4466 | -106.6228
17.4503 | -106.6199
17.454 | -106.6162
17.4577 | -106.6136
17.4614 | -106.6116
17.4651 | -106.6102
17.4688 | -106.6067
17.4725 | -106.6042
17.4761 | -106.6039
17.4798 | -106.6036
17.4835 | -106.5999
17.4872 | -106.5955
17.4909 | -106.5894
17.4946 | -106.5844
17.4983 | -106.5807
17.502 | -106.5775
17.5056 | -106.5752
17.5093 | -106.5737
17.513 | -106.5733
17.5167 | -106.5719
17.5204 | -106.5723
17.5241 | -106.5715
17.5278 | -106.5695
17.5315 | -106.5674
17.5352 | -106.5675
17.5388 | -106.5678
17.5425 | -106.5662
17.5462 | -106.5676
17.5499 | -106.5708
17.5536 | -106.578
17.5573 | -106.5842
17.561 | -106.5901
17.5647 | -106.5968
17.5684 | -106.602
17.572 | -106.6075
17.5757 | -106.6136
17.5794 | -106.6218
17.5831 | -106.6297
17.5868 | -106.6379
17.5905 | -106.6459
17.5942 | -106.6532
17.5979 | -106.6624
17.6015 | -106.6729
17.6052 | -106.6826
17.6089 | -106.6946
17.6126 | -106.7051
17.6163 | -106.7163
17.62 | -106.7265
17.6237 | -106.7332
17.6274 | -106.74
17.6311 | -106.7442
17.6347 | -106.7483
17.6384 | -106.7532
17.6421 | -106.7584
17.6458 | -106.7648
17.6495 | -106.7683
17.6532 | -106.7702
17.6569 | -106.7706
17.6606 | -106.7714
17.6643 | -106.772
17.6679 | -106.7713
17.6716 | -106.7708
17.6753 | -106.7691
17.679 | -106.7651
17.6827 | -106.7619
17.6864 | -106.7604
17.6901 | -106.7591
17.6938 | -106.7567
17.6974 | -106.7486
17.7011 | -106.7411
17.7048 | -106.7361
17.7085 | -106.734
17.7122 | -106.7301
17.7159 | -106.7246
17.7196 | -106.7199
17.7233 | -106.7173
17.727 | -106.7125
17.7306 | -106.7065
17.7343 | -106.705
17.738 | -106.705
17.7417 | -106.7045
17.7454 | -106.7012
17.7491 | -106.6991
17.7528 | -106.6964
17.7565 | -106.6919
17.7602 | -106.6891
17.7638 | -106.6849
17.7675 | -106.6812
17.7712 | -106.6808
17.7749 | -106.6806
17.7786 | -106.6805
17.7823 | -106.6795
17.786 | -106.678
17.7897 | -106.6765
17.7934 | -106.6759
17.797 | -106.6724
17.8007 | -106.6686
17.8044 | -106.666
17.8081 | -106.6639
17.8118 | -106.6626
17.8155 | -106.6575
17.8192 | -106.6524
17.8229 | -106.6496
17.8265 | -106.6458
17.8302 | -106.6432
17.8339 | -106.6389
17.8376 | -106.635
17.8413 | -106.6321
17.845 | -106.6308
17.8487 | -106.6293
17.8524 | -106.6253
17.8561 | -106.6193
17.8597 | -106.6132
17.8634 | -106.607
17.8671 | -106.6023
17.8708 | -106.5979
17.8745 | -106.594
17.8782 | -106.5938
17.8819 | -106.5926
17.8856 | -106.5933
17.8893 | -106.5919
17.8929 | -106.5902
17.8966 | -106.5887
17.9003 | -106.5853
17.904 | -106.5845
17.9077 | -106.5851
17.9114 | -106.584
17.9151 | -106.5814
17.9188 | -106.5804
17.9224 | -106.5775
17.9261 | -106.5743
17.9298 | -106.5734
17.9335 | -106.5712
17.9372 | -106.5685
17.9409 | -106.5677
17.9446 | -106.5638
17.9483 | -106.5594
17.952 | -106.5581
17.9556 | -106.5562
17.9593 | -106.5526
17.963 | -106.547
17.9667 | -106.5398
17.9704 | -106.5331
17.9741 | -106.5277
17.9778 | -106.52
17.9815 | -106.5147
17.9852 | -106.5117
17.9888 | -106.5056
17.9925 | -106.5004
17.9962 | -106.4945
17.9999 | -106.4901
18.0036 | -106.4865
18.0073 | -106.484
18.011 | -106.4831
18.0147 | -106.4797
18.0184 | -106.4789
18.022 | -106.4802
18.0257 | -106.4793
18.0294 | -106.4789
18.0331 | -106.4787
18.0368 | -106.4798
18.0405 | -106.4831
18.0442 | -106.4869
18.0479 | -106.4898
18.0515 | -106.4957
18.0552 | -106.5033
18.0589 | -106.5074
18.0626 | -106.5128
18.0663 | -106.5216
18.07 | -106.5312
18.0737 | -106.5389
18.0774 | -106.5459
18.0811 | -106.5545
18.0847 | -106.5589
18.0884 | -106.5668
18.0921 | -106.5757
18.0958 | -106.584
18.0995 | -106.593
18.1032 | -106.6026
18.1069 | -106.6102
18.1106 | -106.6179
18.1143 | -106.6272
18.1179 | -106.6387
18.1216 | -106.6474
18.1253 | -106.6555
18.129 | -106.662
18.1327 | -106.6679
18.1364 | -106.6716
18.1401 | -106.6786
18.1438 | -106.6871
18.1474 | -106.6949
18.1511 | -106.699
18.1548 | -106.7042
18.1585 | -106.7121
18.1622 | -106.7176
18.1659 | -106.7211
18.1696 | -106.7256
18.1733 | -106.7283
18.177 | -106.7305
18.1806 | -106.7339
18.1843 | -106.7352
18.188 | -106.7326
18.1917 | -106.7307
18.1954 | -106.7304
18.1991 | -106.7268
18.2028 | -106.726
18.2065 | -106.7244
18.2102 | -106.7206
18.2138 | -106.716
18.2175 | -106.7112
18.2212 | -106.7056
18.2249 | -106.7018
18.2286 | -106.6971
18.2323 | -106.6903
18.236 | -106.6833
18.2397 | -106.6755
18.2433 | -106.6688
18.247 | -106.6595
18.2507 | -106.6488
18.2544 | -106.6405
18.2581 | -106.6351
18.2618 | -106.6264
18.2655 | -106.6157
18.2692 | -106.6062
18.2729 | -106.599
18.2765 | -106.5898
18.2802 | -106.5802
18.2839 | -106.5726
18.2876 | -106.5659
18.2913 | -106.5596
18.295 | -106.5533
18.2987 | -106.5473
18.3024 | -106.5434
18.3061 | -106.5398
18.3097 | -106.5337
18.3134 | -106.5279
18.3171 | -106.5241
18.3208 | -106.5209
18.3245 | -106.5187
18.3282 | -106.5153
18.3319 | -106.5137
18.3356 | -106.5129
18.3392 | -106.51
18.3429 | -106.5108
18.3466 | -106.5108
18.3503 | -106.5092
18.354 | -106.5101
18.3577 | -106.5114
18.3614 | -106.5117
18.3651 | -106.5112
18.3688 | -106.5122
18.3724 | -106.5135
18.3761 | -106.5145
18.3798 | -106.5139
18.3835 | -106.5158
18.3872 | -106.5167
18.3909 | -106.5199
18.3946 | -106.5219
18.3983 | -106.5229
18.402 | -106.5222
18.4056 | -106.5205
18.4093 | -106.5206
18.413 | -106.522
18.4167 | -106.5216
18.4204 | -106.5212
18.4241 | -106.5215
18.4278 | -106.523
18.4315 | -106.5222
18.4352 | -106.5232
18.4388 | -106.5242
18.4425 | -106.5216
18.4462 | -106.5197
18.4499 | -106.5195
18.4536 | -106.5194
18.4573 | -106.5164
18.461 | -106.5122
18.4647 | -106.5075
18.4683 | -106.5061
18.472 | -106.5037
18.4757 | -106.4991
18.4794 | -106.4917
18.4831 | -106.4855
18.4868 | -106.4802
18.4905 | -106.4752
18.4942 | -106.4711
18.4979 | -106.4673
18.5015 | -106.4637
18.5052 | -106.4561
18.5089 | -106.4505
18.5126 | -106.4451
18.5163 | -106.4373
18.52 | -106.427
18.5237 | -106.4191
18.5274 | -106.4149
18.5311 | -106.4097
18.5347 | -106.4021
18.5384 | -106.3969
18.5421 | -106.3919
18.5458 | -106.3866
18.5495 | -106.3803
18.5532 | -106.3738
18.5569 | -106.3677
18.5606 | -106.3624
18.5642 | -106.3571
18.5679 | -106.3564
18.5716 | -106.3568
18.5753 | -106.3574
18.579 | -106.358
18.5827 | -106.3574
18.5864 | -106.3563
18.5901 | -106.3579
18.5938 | -106.3584
18.5974 | -106.36
18.6011 | -106.3647
18.6048 | -106.3694
18.6085 | -106.3724
18.6122 | -106.3735
18.6159 | -106.3742
18.6196 | -106.3763
18.6233 | -106.3787
18.627 | -106.3823
18.6306 | -106.3865
18.6343 | -106.3904
18.638 | -106.3942
18.6417 | -106.3974
18.6454 | -106.4004
18.6491 | -106.4004
18.6528 | -106.4016
18.6565 | -106.3999
18.6602 | -106.3999
18.6638 | -106.3997
18.6675 | -106.3971
18.6712 | -106.3949
18.6749 | -106.3921
18.6786 | -106.3887
18.6823 | -106.3865
18.686 | -106.3829
18.6897 | -106.379
18.6933 | -106.3754
18.697 | -106.3738
18.7007 | -106.3713
18.7044 | -106.3685
18.7081 | -106.3694
18.7118 | -106.37
18.7155 | -106.369
18.7192 | -106.3686
18.7229 | -106.3697
18.7265 | -106.3729
18.7302 | -106.3752
18.7339 | -106.3785
18.7376 | -106.3838
18.7413 | -106.3922
18.745 | -106.4004
18.7487 | -106.4077
18.7524 | -106.4158
18.7561 | -106.4248
18.7597 | -106.4333
18.7634 | -106.4431
18.7671 | -106.451
18.7708 | -106.4621
18.7745 | -106.4754
18.7782 | -106.4885
18.7819 | -106.5022
18.7856 | -106.5142
18.7892 | -106.5267
18.7929 | -106.5386
18.7966 | -106.5506
18.8003 | -106.5638
18.804 | -106.5748
18.8077 | -106.5852
18.8114 | -106.5963
18.8151 | -106.6074
18.8188 | -106.6197
18.8224 | -106.6324
18.8261 | -106.6443
18.8298 | -106.6549
18.8335 | -106.6614
18.8372 | -106.6707
18.8409 | -106.6806
18.8446 | -106.6894
18.8483 | -106.7
18.852 | -106.71
18.8556 | -106.7202
18.8593 | -106.7317
18.863 | -106.7412
18.8667 | -106.7491
18.8704 | -106.7566
18.8741 | -106.7642
18.8778 | -106.7759
18.8815 | -106.7857
18.8851 | -106.7933
18.8888 | -106.8006
18.8925 | -106.8076
18.8962 | -106.8119
18.8999 | -106.8164
18.9036 | -106.8225
18.9073 | -106.8285
18.911 | -106.8345
18.9147 | -106.8409
18.9183 | -106.8484
18.922 | -106.8554
18.9257 | -106.8631
18.9294 | -106.8694
18.9331 | -106.8737
18.9368 | -106.8794
18.9405 | -106.8852
18.9442 | -106.8905
18.9479 | -106.897
18.9515 | -106.9025
18.9552 | -106.9075
18.9589 | -106.9096
18.9626 | -106.9106
18.9663 | -106.9099
18.97 | -106.9085
18.9737 | -106.905
18.9774 | -106.9032
18.981 | -106.9016
18.9847 | -106.9009
18.9884 | -106.8986
18.9921 | -106.8966
18.9958 | -106.8932
18.9995 | -106.8885
19.0032 | -106.8809
19.0069 | -106.8728
19.0106 | -106.8627
19.0142 | -106.8577
19.0179 | -106.852
19.0216 | -106.845
19.0253 | -106.8359
19.029 | -106.8283
19.0327 | -106.8201
19.0364 | -106.8126
19.0401 | -106.8038
19.0438 | -106.7924
19.0474 | -106.7819
19.0511 | -106.7739
19.0548 | -106.769
19.0585 | -106.7611
19.0622 | -106.7528
19.0659 | -106.7464
19.0696 | -106.7415
19.0733 | -106.7375
19.077 | -106.7331
19.0806 | -106.7276
19.0843 | -106.7228
19.088 | -106.7156
19.0917 | -106.7074
19.0954 | -106.6986
19.0991 | -106.6941
19.1028 | -106.6912
19.1065 | -106.6856
19.1101 | -106.678
19.1138 | -106.6722
19.1175 | -106.6664
19.1212 | -106.661
19.1249 | -106.6553
19.1286 | -106.6496
19.1323 | -106.6444
19.136 | -106.6372
19.1397 | -106.6302
19.1433 | -106.6273
19.147 | -106.6234
19.1507 | -106.6189
19.1544 | -106.6178
19.1581 | -106.6135
19.1618 | -106.6098
19.1655 | -106.6076
19.1692 | -106.6078
19.1729 | -106.6084
19.1765 | -106.6075
19.1802 | -106.6081
19.1839 | -106.608
19.1876 | -106.6075
19.1913 | -106.6087
19.195 | -106.6089
19.1987 | -106.6098
19.2024 | -106.6085
19.206 | -106.6105
19.2097 | -106.6142
19.2134 | -106.6161
19.2171 | -106.6184
19.2208 | -106.6214
19.2245 | -106.6238
19.2282 | -106.627
19.2319 | -106.634
19.2356 | -106.6415
19.2392 | -106.6474
19.2429 | -106.6522
19.2466 | -106.6572
19.2503 | -106.66
19.254 | -106.6626
19.2577 | -106.6656
19.2614 | -106.6686
19.2651 | -106.6713
19.2688 | -106.6745
19.2724 | -106.6777
19.2761 | -106.679
19.2798 | -106.681
19.2835 | -106.6844
19.2872 | -106.6827
19.2909 | -106.6803
19.2946 | -106.6781
19.2983 | -106.6785
19.302 | -106.6765
19.3056 | -106.6725
19.3093 | -106.6685
19.313 | -106.664
19.3167 | -106.6588
19.3204 | -106.6546
19.3241 | -106.647
19.3278 | -106.6377
19.3315 | -106.6316
19.3351 | -106.6239
19.3388 | -106.6144
19.3425 | -106.6028
19.3462 | -106.5932
19.3499 | -106.5848
19.3536 | -106.5747
19.3573 | -106.5651
19.361 | -106.5588
19.3647 | -106.5511
19.3683 | -106.5419
19.372 | -106.5306
19.3757 | -106.5198
19.3794 | -106.5114
19.3831 | -106.504
19.3868 | -106.4941
19.3905 | -106.4855
19.3942 | -106.4787
19.3979 | -106.4726
19.4015 | -106.4664
19.4052 | -106.461
19.4089 | -106.4569
19.4126 | -106.4535
19.4163 | -106.4494
19.42 | -106.4458
19.4237 | -106.4456
19.4274 | -106.4468
19.431 | -106.4468
19.4347 | -106.4499
19.4384 | -106.4524
19.4421 | -106.4542
19.4458 | -106.4578
19.4495 | -106.4605
19.4532 | -106.465
19.4569 | -106.4688
19.4606 | -106.4734
19.4642 | -106.4782
19.4679 | -106.4808
19.4716 | -106.4851
19.4753 | -106.49
19.479 | -106.4941
19.4827 | -106.4965
19.4864 | -106.4985
19.4901 | -106.4999
19.4938 | -106.5013
19.4974 | -106.502
19.5011 | -106.5038
19.5048 | -106.5042
19.5085 | -106.5013
19.5122 | -106.496
19.5159 | -106.4922
19.5196 | -106.4905
19.5233 | -106.4878
19.5269 | -106.4836
19.5306 | -106.4801
19.5343 | -106.4766
19.538 | -106.4728
19.5417 | -106.4695
19.5454 | -106.4669
19.5491 | -106.4639
19.5528 | -106.4622
19.5565 | -106.4597
19.5601 | -106.4587
19.5638 | -106.4579
19.5675 | -106.4573
19.5712 | -106.457
19.5749 | -106.4574
19.5786 | -106.4572
19.5823 | -106.4574
19.586 | -106.4602
19.5897 | -106.4647
19.5933 | -106.4699
19.597 | -106.4727
19.6007 | -106.4772
19.6044 | -106.4836
19.6081 | -106.4889
19.6118 | -106.4943
19.6155 | -106.501
19.6192 | -106.5106
19.6228 | -106.5187
19.6265 | -106.5288
19.6302 | -106.5416
19.6339 | -106.5544
19.6376 | -106.5657
19.6413 | -106.5794
19.645 | -106.5948
19.6487 | -106.6093
19.6524 | -106.6236
19.656 | -106.6382
19.6597 | -106.6538
19.6634 | -106.672
19.6671 | -106.6885
19.6708 | -106.7026
19.6745 | -106.7152
19.6782 | -106.7294
19.6819 | -106.7421
19.6856 | -106.749
19.6892 | -106.7629
19.6929 | -106.7759
19.6966 | -106.7863
19.7003 | -106.7943
19.704 | -106.7996
19.7077 | -106.8034
19.7114 | -106.8084
19.7151 | -106.8134
19.7188 | -106.8199
19.7224 | -106.8241
19.7261 | -106.8271
19.7298 | -106.8305
19.7335 | -106.8302
19.7372 | -106.8296
19.7409 | -106.8301
19.7446 | -106.8318
19.7483 | -106.8321
19.7519 | -106.8297
19.7556 | -106.8278
19.7593 | -106.8261
19.763 | -106.8246
19.7667 | -106.8209
19.7704 | -106.8153
19.7741 | -106.8091
19.7778 | -106.8042
19.7815 | -106.7982
19.7851 | -106.7912
19.7888 | -106.7863
19.7925 | -106.7828
19.7962 | -106.7785
19.7999 | -106.7723
19.8036 | -106.7674
19.8073 | -106.7613
19.811 | -106.7547
19.8147 | -106.7479
19.8183 | -106.7417
19.822 | -106.7344
19.8257 | -106.7268
19.8294 | -106.7198
19.8331 | -106.7136
19.8368 | -106.7082
19.8405 | -106.7007
19.8442 | -106.6939
19.8478 | -106.6883
19.8515 | -106.6815
19.8552 | -106.6742
19.8589 | -106.667
19.8626 | -106.6606
19.8663 | -106.6538
19.87 | -106.6475
19.8737 | -106.6427
19.8774 | -106.6384
19.881 | -106.6353
19.8847 | -106.6299
19.8884 | -106.6228
19.8921 | -106.6205
19.8958 | -106.6186
19.8995 | -106.617
19.9032 | -106.6158
19.9069 | -106.6137
19.9106 | -106.611
19.9142 | -106.6091
19.9179 | -106.609
19.9216 | -106.6077
19.9253 | -106.6067
19.929 | -106.6059
19.9327 | -106.6062
19.9364 | -106.6069
19.9401 | -106.6083
19.9438 | -106.6091
19.9474 | -106.6086
19.9511 | -106.6091
19.9548 | -106.6101
19.9585 | -106.6101
19.9622 | -106.6117
19.9659 | -106.6123
19.9696 | -106.6107
19.9733 | -106.6094
19.9769 | -106.6093
19.9806 | -106.6088
19.9843 | -106.6092
19.988 | -106.6098
19.9917 | -106.6096
19.9954 | -106.6094
19.9991 | -106.608
20.0028 | -106.6049
20.0065 | -106.6029
20.0101 | -106.6018
20.0138 | -106.6021
20.0175 | -106.6047
20.0212 | -106.6065
20.0249 | -106.6057
20.0286 | -106.6042
20.0323 | -106.6035
20.036 | -106.6026
20.0397 | -106.6004
20.0433 | -106.6022
20.047 | -106.6044
20.0507 | -106.6068
20.0544 | -106.61
20.0581 | -106.6133
20.0618 | -106.6146
20.0655 | -106.6171
20.0692 | -106.6209
20.0728 | -106.6236
20.0765 | -106.6263
20.0802 | -106.6296
20.0839 | -106.6329
20.0876 | -106.6347
20.0913 | -106.6396
20.095 | -106.6445
20.0987 | -106.6488
20.1024 | -106.6538
20.106 | -106.6595
20.1097 | -106.6641
20.1134 | -106.6702
20.1171 | -106.6774
20.1208 | -106.6839
20.1245 | -106.6894
20.1282 | -106.695
20.1319 | -106.7023
20.1356 | -106.7103
20.1392 | -106.7174
20.1429 | -106.7264
20.1466 | -106.7343
20.1503 | -106.7415
20.154 | -106.7486
20.1577 | -106.7554
20.1614 | -106.7616
20.1651 | -106.7665
20.1687 | -106.7714
20.1724 | -106.7788
20.1761 | -106.7844
20.1798 | -106.7902
20.1835 | -106.7947
20.1872 | -106.8003
20.1909 | -106.8059
20.1946 | -106.8122
20.1983 | -106.8163
20.2019 | -106.8193
20.2056 | -106.8247
20.2093 | -106.8283
20.213 | -106.8331
20.2167 | -106.8375
20.2204 | -106.8389
20.2241 | -106.8388
20.2278 | -106.8395
20.2315 | -106.8402
20.2351 | -106.8418
20.2388 | -106.8434
20.2425 | -106.8434
20.2462 | -106.844
20.2499 | -106.8428
20.2536 | -106.8411
20.2573 | -106.839
20.261 | -106.8354
20.2646 | -106.8299
20.2683 | -106.8257
20.272 | -106.8225
20.2757 | -106.8195
20.2794 | -106.8173
20.2831 | -106.8131
20.2868 | -106.8088
20.2905 | -106.8051
20.2942 | -106.8012
20.2978 | -106.7956
20.3015 | -106.7903
20.3052 | -106.7858
20.3089 | -106.7815
20.3126 | -106.7772
20.3163 | -106.7717
20.32 | -106.7681
20.3237 | -106.7663
20.3274 | -106.7631
20.331 | -106.7574
20.3347 | -106.75
20.3384 | -106.7442
20.3421 | -106.7403
20.3458 | -106.7355
20.3495 | -106.7318
20.3532 | -106.7279
20.3569 | -106.7233
20.3606 | -106.7196
20.3642 | -106.7153
20.3679 | -106.7098
20.3716 | -106.7065
20.3753 | -106.7046
20.379 | -106.7026
20.3827 | -106.6985
20.3864 | -106.6942
20.3901 | -106.6895
20.3937 | -106.6856
20.3974 | -106.6835
20.4011 | -106.6813
20.4048 | -106.6776
20.4085 | -106.6738
20.4122 | -106.6703
20.4159 | -106.6654
20.4196 | -106.6592
20.4233 | -106.6553
20.4269 | -106.6527
20.4306 | -106.648
20.4343 | -106.6435
20.438 | -106.6378
20.4417 | -106.6325
20.4454 | -106.626
20.4491 | -106.621
20.4528 | -106.6153
20.4565 | -106.6108
20.4601 | -106.6049
20.4638 | -106.5969
20.4675 | -106.589
20.4712 | -106.5808
20.4749 | -106.5732
20.4786 | -106.5643
20.4823 | -106.5555
20.486 | -106.5492
20.4896 | -106.5416
20.4933 | -106.5325
20.497 | -106.5253
20.5007 | -106.5194
20.5044 | -106.5137
20.5081 | -106.5085
20.5118 | -106.5051
20.5155 | -106.503
20.5192 | -106.5002
20.5228 | -106.4984
20.5265 | -106.4955
20.5302 | -106.4899
20.5339 | -106.4862
20.5376 | -106.4822
20.5413 | -106.4802
20.545 | -106.4787
20.5487 | -106.4785
20.5524 | -106.4784
20.556 | -106.4791
20.5597 | -106.4789
20.5634 | -106.478
20.5671 | -106.4769
20.5708 | -106.4766
20.5745 | -106.478
20.5782 | -106.4805
20.5819 | -106.4836
20.5856 | -106.4858
20.5892 | -106.4885
20.5929 | -106.4908
20.5966 | -106.4913
20.6003 | -106.4915
20.604 | -106.4932
20.6077 | -106.4947
20.6114 | -106.4962
20.6151 | -106.4948
20.6187 | -106.495
20.6224 | -106.4942
20.6261 | -106.4944
20.6298 | -106.4938
20.6335 | -106.4918
20.6372 | -106.4888
20.6409 | -106.4856
20.6446 | -106.4823
20.6483 | -106.4796
20.6519 | -106.4744
20.6556 | -106.4692
20.6593 | -106.4642
20.663 | -106.4587
20.6667 | -106.4519
20.6704 | -106.4445
20.6741 | -106.4411
20.6778 | -106.4372
20.6815 | -106.4323
20.6851 | -106.427
20.6888 | -106.4211
20.6925 | -106.4143
20.6962 | -106.4102
20.6999 | -106.4076
20.7036 | -106.4067
20.7073 | -106.4052
20.711 | -106.4038
20.7146 | -106.4004
20.7183 | -106.3973
20.722 | -106.397
20.7257 | -106.3952
20.7294 | -106.3931
20.7331 | -106.3912
20.7368 | -106.3904
20.7405 | -106.3903
20.7442 | -106.3911
20.7478 | -106.3905
20.7515 | -106.3865
20.7552 | -106.3856
20.7589 | -106.3872
20.7626 | -106.387
20.7663 | -106.3868
20.77 | -106.3847
20.7737 | -106.3843
20.7774 | -106.3832
20.781 | -106.3803
20.7847 | -106.3776
20.7884 | -106.3753
20.7921 | -106.3733
20.7958 | -106.3691
20.7995 | -106.3644
20.8032 | -106.3597
20.8069 | -106.3541
20.8105 | -106.3482
20.8142 | -106.342
20.8179 | -106.3378
20.8216 | -106.3354
20.8253 | -106.3351
20.829 | -106.3345
20.8327 | -106.3338
20.8364 | -106.3323
20.8401 | -106.3308
20.8437 | -106.3299
20.8474 | -106.3308
20.8511 | -106.33
20.8548 | -106.3315
20.8585 | -106.3342
20.8622 | -106.3363
20.8659 | -106.3377
20.8696 | -106.3408
20.8733 | -106.3431
20.8769 | -106.3476
20.8806 | -106.3553
20.8843 | -106.3622
20.888 | -106.3698
20.8917 | -106.3782
20.8954 | -106.3893
20.8991 | -106.4004
20.9028 | -106.4119
20.9064 | -106.4225
20.9101 | -106.435
20.9138 | -106.4486
20.9175 | -106.4639
20.9212 | -106.4772
20.9249 | -106.4899
20.9286 | -106.5012
20.9323 | -106.5135
20.936 | -106.5255
20.9396 | -106.5368
20.9433 | -106.5476
20.947 | -106.559
20.9507 | -106.5706
20.9544 | -106.5799
20.9581 | -106.5867
20.9618 | -106.5955
20.9655 | -106.6038
20.9692 | -106.6108
20.9728 | -106.6184
20.9765 | -106.6241
20.9802 | -106.6298
20.9839 | -106.6343
20.9876 | -106.6385
20.9913 | -106.6402
20.995 | -106.6416
20.9987 | -106.6412
21.0024 | -106.6417
21.006 | -106.6417
21.0097 | -106.6417
21.0134 | -106.6408
21.0171 | -106.6398
21.0208 | -106.6401
21.0245 | -106.6408
21.0282 | -106.6385
21.0319 | -106.636
21.0355 | -106.634
21.0392 | -106.6302
21.0429 | -106.6277
21.0466 | -106.6258
21.0503 | -106.625
21.054 | -106.6244
21.0577 | -106.6212
21.0614 | -106.6189
21.0651 | -106.6168
21.0687 | -106.6145
21.0724 | -106.6128
21.0761 | -106.6116
21.0798 | -106.6113
21.0835 | -106.6118
21.0872 | -106.6125
21.0909 | -106.6127
21.0946 | -106.6122
21.0983 | -106.613
21.1019 | -106.6144
21.1056 | -106.6148
21.1093 | -106.6157
21.113 | -106.6169
21.1167 | -106.6174
21.1204 | -106.6173
21.1241 | -106.6177
21.1278 | -106.6172
21.1314 | -106.6172
21.1351 | -106.6191
21.1388 | -106.6199
21.1425 | -106.6202
21.1462 | -106.6178
21.1499 | -106.6162
21.1536 | -106.6142
21.1573 | -106.6139
21.161 | -106.6158
21.1646 | -106.6194
21.1683 | -106.6209
21.172 | -106.623
21.1757 | -106.6236
21.1794 | -106.6217
21.1831 | -106.6199
21.1868 | -106.6173
21.1905 | -106.6173
21.1942 | -106.6126
21.1978 | -106.6089
21.2015 | -106.6067
21.2052 | -106.6031
21.2089 | -106.5992
21.2126 | -106.5948
21.2163 | -106.5908
21.22 | -106.5886
21.2237 | -106.5868
21.2274 | -106.5854
21.231 | -106.5819
21.2347 | -106.5764
21.2384 | -106.5702
21.2421 | -106.5661
21.2458 | -106.562
21.2495 | -106.5546
21.2532 | -106.5501
21.2569 | -106.5454
21.2605 | -106.5374
21.2642 | -106.5268
21.2679 | -106.5169
21.2716 | -106.509
21.2753 | -106.5014
21.279 | -106.4938
21.2827 | -106.4881
21.2864 | -106.4806
21.2901 | -106.4737
21.2937 | -106.4672
21.2974 | -106.4597
21.3011 | -106.4506
21.3048 | -106.4426
21.3085 | -106.4355
21.3122 | -106.4279
21.3159 | -106.4205
21.3196 | -106.4126
21.3233 | -106.4075
21.3269 | -106.4012
21.3306 | -106.3965
21.3343 | -106.393
21.338 | -106.3875
21.3417 | -106.384
21.3454 | -106.3814
21.3491 | -106.379
21.3528 | -106.3744
21.3564 | -106.3705
21.3601 | -106.3673
21.3638 | -106.3656
21.3675 | -106.3654
21.3712 | -106.3658
21.3749 | -106.3624
21.3786 | -106.3592
21.3823 | -106.358
21.386 | -106.3583
21.3896 | -106.3581
21.3933 | -106.357
21.397 | -106.3576
21.4007 | -106.357
21.4044 | -106.356
21.4081 | -106.3584
21.4118 | -106.3602
21.4155 | -106.3604
21.4192 | -106.3597
21.4228 | -106.36
21.4265 | -106.3599
21.4302 | -106.3611
21.4339 | -106.3617
21.4376 | -106.3584
21.4413 | -106.3574
21.445 | -106.3578
21.4487 | -106.358
21.4523 | -106.3597
21.456 | -106.3607
21.4597 | -106.3605
21.4634 | -106.3596
21.4671 | -106.3581
21.4708 | -106.3563
21.4745 | -106.3562
21.4782 | -106.3555
21.4819 | -106.3563
21.4855 | -106.3554
21.4892 | -106.353
21.4929 | -106.3513
21.4966 | -106.3522
21.5003 | -106.3514
21.504 | -106.3494
21.5077 | -106.348
21.5114 | -106.3454
21.5151 | -106.3433
21.5187 | -106.3405
21.5224 | -106.3377
21.5261 | -106.3346
21.5298 | -106.3319
21.5335 | -106.3301
21.5372 | -106.3291
21.5409 | -106.3263
21.5446 | -106.3243
21.5482 | -106.3229
21.5519 | -106.3227
21.5556 | -106.3206
21.5593 | -106.3212
21.563 | -106.3251
21.5667 | -106.3305
21.5704 | -106.3339
21.5741 | -106.3379
21.5778 | -106.3421
21.5814 | -106.3451
21.5851 | -106.349
21.5888 | -106.3515
21.5925 | -106.3568
21.5962 | -106.3633
21.5999 | -106.369
21.6036 | -106.3753
21.6073 | -106.3856
21.611 | -106.3944
21.6146 | -106.4002
21.6183 | -106.4074
21.622 | -106.4159
21.6257 | -106.4249
21.6294 | -106.4306
21.6331 | -106.4356
21.6368 | -106.4397
21.6405 | -106.4432
21.6442 | -106.4471
21.6478 | -106.4519
21.6515 | -106.457
21.6552 | -106.4612
21.6589 | -106.4652
21.6626 | -106.4678
21.6663 | -106.4712
21.67 | -106.4741
21.6737 | -106.476
21.6773 | -106.4746
21.681 | -106.4762
21.6847 | -106.479
21.6884 | -106.4782
21.6921 | -106.4751
21.6958 | -106.4716
21.6995 | -106.4703
21.7032 | -106.47
21.7069 | -106.4693
21.7105 | -106.4676
21.7142 | -106.4651
21.7179 | -106.4638
21.7216 | -106.4625
21.7253 | -106.4606
21.729 | -106.4586
21.7327 | -106.4565
21.7364 | -106.4548
21.7401 | -106.4536
21.7437 | -106.4558
21.7474 | -106.4562
21.7511 | -106.4545
21.7548 | -106.4552
21.7585 | -106.4538
21.7622 | -106.4523
21.7659 | -106.4505
21.7696 | -106.4502
21.7732 | -106.4518
21.7769 | -106.455
21.7806 | -106.4556
21.7843 | -106.4564
21.788 | -106.4584
21.7917 | -106.4632
21.7954 | -106.4679
21.7991 | -106.4709
21.8028 | -106.4763
21.8064 | -106.4795
21.8101 | -106.4818
21.8138 | -106.4818
21.8175 | -106.4805
21.8212 | -106.4821
21.8249 | -106.4841
21.8286 | -106.4867
21.8323 | -106.4888
21.836 | -106.4909
21.8396 | -106.4932
21.8433 | -106.4959
21.847 | -106.4968
21.8507 | -106.4961
21.8544 | -106.4915
21.8581 | -106.4878
21.8618 | -106.4839
21.8655 | -106.4807
21.8692 | -106.4744
21.8728 | -106.4659
21.8765 | -106.4572
21.8802 | -106.4485
21.8839 | -106.437
21.8876 | -106.4251
21.8913 | -106.4126
21.895 | -106.3992
21.8987 | -106.3868
21.9023 | -106.3723
21.906 | -106.3561
21.9097 | -106.3399
21.9134 | -106.3236
21.9171 | -106.3067
21.9208 | -106.2892
21.9245 | -106.2731
21.9282 | -106.2565
21.9319 | -106.2383
21.9355 | -106.2183
21.9392 | -106.1999
21.9429 | -106.1826
21.9466 | -106.1662
21.9503 | -106.1486
21.954 | -106.1306
21.9577 | -106.1121
21.9614 | -106.0922
21.9651 | -106.0749
21.9687 | -106.0587
21.9724 | -106.0419
21.9761 | -106.0251
21.9798 | -106.0087
21.9835 | -105.9937
21.9872 | -105.9813
21.9909 | -105.9691
21.9946 | -105.9554
21.9982 | -105.9408
22.0019 | -105.928
22.0056 | -105.9177
22.0093 | -105.9067
22.013 | -105.8953
22.0167 | -105.8858
22.0204 | -105.8783
22.0241 | -105.8723
22.0278 | -105.8654
22.0314 | -105.8598
22.0351 | -105.8556
22.0388 | -105.8529
22.0425 | -105.852
22.0462 | -105.8504
22.0499 | -105.8488
22.0536 | -105.8489
22.0573 | -105.8508
22.061 | -105.8552
22.0646 | -105.8601
22.0683 | -105.8658
22.072 | -105.8713
22.0757 | -105.877
22.0794 | -105.8824
22.0831 | -105.889
22.0868 | -105.8967
22.0905 | -105.9063
22.0941 | -105.9171
22.0978 | -105.9293
22.1015 | -105.9404
22.1052 | -105.9523
22.1089 | -105.9629
22.1126 | -105.9728
22.1163 | -105.9853
22.12 | -105.9993
22.1237 | -106.0153
22.1273 | -106.028
22.131 | -106.0406
22.1347 | -106.0544
22.1384 | -106.068
22.1421 | -106.0803
22.1458 | -106.0934
22.1495 | -106.1067
22.1532 | -106.1204
22.1569 | -106.1317
22.1605 | -106.144
22.1642 | -106.1591
22.1679 | -106.1734
22.1716 | -106.1846
22.1753 | -106.1962
22.179 | -106.2057
22.1827 | -106.2144
22.1864 | -106.223
22.19 | -106.2297
22.1937 | -106.2358
22.1974 | -106.2405
22.2011 | -106.2462
22.2048 | -106.2517
22.2085 | -106.2545
22.2122 | -106.2553
22.2159 | -106.2568
22.2196 | -106.2577
22.2232 | -106.2576
22.2269 | -106.2555
22.2306 | -106.2517
22.2343 | -106.2493
22.238 | -106.246
22.2417 | -106.2403
22.2454 | -106.2344
22.2491 | -106.2278
22.2528 | -106.2217
22.2564 | -106.2182
22.2601 | -106.2141
22.2638 | -106.2091
22.2675 | -106.2041
22.2712 | -106.1997
22.2749 | -106.1953
22.2786 | -106.1909
22.2823 | -106.1863
22.286 | -106.1832
22.2896 | -106.1792
22.2933 | -106.1769
22.297 | -106.1756
22.3007 | -106.1739
22.3044 | -106.1716
22.3081 | -106.1705
22.3118 | -106.1712
22.3155 | -106.1705
22.3191 | -106.1703
22.3228 | -106.1712
22.3265 | -106.1722
22.3302 | -106.1715
22.3339 | -106.1717
22.3376 | -106.17
22.3413 | -106.1695
22.345 | -106.1716
22.3487 | -106.1745
22.3523 | -106.1753
22.356 | -106.1759
22.3597 | -106.1758
22.3634 | -106.177
22.3671 | -106.1778
22.3708 | -106.1784
22.3745 | -106.1774
22.3782 | -106.1742
22.3819 | -106.1714
22.3855 | -106.1705
22.3892 | -106.1696
22.3929 | -106.1696
22.3966 | -106.169
22.4003 | -106.1675
22.404 | -106.1652
22.4077 | -106.1617
22.4114 | -106.1586
22.415 | -106.1553
22.4187 | -106.153
22.4224 | -106.1509
22.4261 | -106.1477
22.4298 | -106.1429
22.4335 | -106.1399
22.4372 | -106.1349
22.4409 | -106.1289
22.4446 | -106.1258
22.4482 | -106.1227
22.4519 | -106.1186
22.4556 | -106.1128
22.4593 | -106.107
22.463 | -106.1023
22.4667 | -106.0994
22.4704 | -106.0953
22.4741 | -106.0921
22.4778 | -106.0875
22.4814 | -106.0839
22.4851 | -106.08
22.4888 | -106.0768
22.4925 | -106.0737
22.4962 | -106.0719
22.4999 | -106.0701
22.5036 | -106.0682
22.5073 | -106.0665
22.511 | -106.0653
22.5146 | -106.0659
22.5183 | -106.0635
22.522 | -106.0607
22.5257 | -106.0575
22.5294 | -106.0546
22.5331 | -106.0499
22.5368 | -106.0452
22.5405 | -106.0395
22.5441 | -106.0322
22.5478 | -106.0226
22.5515 | -106.0064
22.5552 | -105.9852
22.5589 | -105.9649
22.5626 | -105.9416
22.5663 | -105.9166
22.57 | -105.886
22.5737 | -105.8496
22.5773 | -105.8111
22.581 | -105.7746
22.5847 | -105.7371
22.5884 | -105.6971
22.5921 | -105.6565
22.5958 | -105.6095
22.5995 | -105.5628
22.6032 | -105.5141
22.6069 | -105.4644
22.6105 | -105.409
22.6142 | -105.352
22.6179 | -105.2938
22.6216 | -105.2351
22.6253 | -105.1746
22.629 | -105.111
22.6327 | -105.0449
22.6364 | -104.9774
22.64 | -104.9101
22.6437 | -104.8413
22.6474 | -104.7733
22.6511 | -104.7045
22.6548 | -104.6361
22.6585 | -104.5685
22.6622 | -104.4999
22.6659 | -104.4325
22.6696 | -104.3675
22.6732 | -104.3046
22.6769 | -104.2428
22.6806 | -104.1838
22.6843 | -104.1292
22.688 | -104.0779
22.6917 | -104.0286
22.6954 | -103.9831
22.6991 | -103.9427
22.7028 | -103.9063
22.7064 | -103.8729
22.7101 | -103.8424
22.7138 | -103.8168
22.7175 | -103.796
22.7212 | -103.7801
22.7249 | -103.7697
22.7286 | -103.7648
22.7323 | -103.7643
22.7359 | -103.7677
22.7396 | -103.7762
22.7433 | -103.7898
22.747 | -103.808
22.7507 | -103.8315
22.7544 | -103.8579
22.7581 | -103.8898
22.7618 | -103.9261
22.7655 | -103.9649
22.7691 | -104.0076
22.7728 | -104.0532
22.7765 | -104.1023
22.7802 | -104.1549
22.7839 | -104.2095
22.7876 | -104.266
22.7913 | -104.3257
22.795 | -104.3867
22.7987 | -104.4496
22.8023 | -104.5126
22.806 | -104.5776
22.8097 | -104.6429
22.8134 | -104.7075
22.8171 | -104.7717
22.8208 | -104.835
22.8245 | -104.8985
22.8282 | -104.9611
22.8318 | -105.0216
22.8355 | -105.0806
22.8392 | -105.1394
22.8429 | -105.1961
22.8466 | -105.2518
22.8503 | -105.304
22.854 | -105.3522
22.8577 | -105.3998
22.8614 | -105.4463
22.865 | -105.4927
22.8687 | -105.538
22.8724 | -105.5838
22.8761 | -105.6296
22.8798 | -105.6405
22.8835 | -105.6782
22.8872 | -105.7126
22.8909 | -105.7447
22.8946 | -105.7765
22.8982 | -105.8053
22.9019 | -105.829
22.9056 | -105.845
22.9093 | -105.8574
22.913 | -105.8686
22.9167 | -105.877
22.9204 | -105.8853
22.9241 | -105.8892
22.9278 | -105.8931
22.9314 | -105.8945
22.9351 | -105.8961
22.9388 | -105.8947
22.9425 | -105.8924
22.9462 | -105.89
22.9499 | -105.8871
22.9536 | -105.8837
22.9573 | -105.8803
22.9609 | -105.8762
22.9646 | -105.8702
22.9683 | -105.8634
22.972 | -105.8557
22.9757 | -105.848
22.9794 | -105.8398
22.9831 | -105.8317
22.9868 | -105.8245
22.9905 | -105.8156
22.9941 | -105.8048
22.9978 | -105.7952
23.0015 | -105.7869
23.0052 | -105.7785
23.0089 | -105.7684
23.0126 | -105.7585
23.0163 | -105.7509
23.02 | -105.7424
23.0237 | -105.7327
23.0273 | -105.7221
23.031 | -105.7099
23.0347 | -105.7017
23.0384 | -105.6913
23.0421 | -105.683
23.0458 | -105.6728
23.0495 | -105.6615
23.0532 | -105.6511
23.0568 | -105.6419
23.0605 | -105.634
23.0642 | -105.6264
23.0679 | -105.6191
23.0716 | -105.6134
23.0753 | -105.6062
23.079 | -105.5987
23.0827 | -105.5918
23.0864 | -105.585
23.09 | -105.5798
23.0937 | -105.5744
23.0974 | -105.5712
23.1011 | -105.568
23.1048 | -105.564
23.1085 | -105.5603
23.1122 | -105.5557
23.1159 | -105.5517
23.1196 | -105.5493
23.1232 | -105.547
23.1269 | -105.5456
23.1306 | -105.5419
23.1343 | -105.5381
23.138 | -105.5376
23.1417 | -105.5378
23.1454 | -105.539
23.1491 | -105.5389
23.1528 | -105.5405
23.1564 | -105.5426
23.1601 | -105.5433
23.1638 | -105.5431
23.1675 | -105.5434
23.1712 | -105.5455
23.1749 | -105.5489
23.1786 | -105.5526
23.1823 | -105.5554
23.1859 | -105.5578
23.1896 | -105.5588
23.1933 | -105.5608
23.197 | -105.5637
23.2007 | -105.567
23.2044 | -105.5715
23.2081 | -105.5764
23.2118 | -105.5821
23.2155 | -105.5885
23.2191 | -105.5937
23.2228 | -105.5967
23.2265 | -105.6014
23.2302 | -105.6078
23.2339 | -105.6139
23.2376 | -105.6211
23.2413 | -105.6281
23.245 | -105.634
23.2487 | -105.6409
23.2523 | -105.6484
23.256 | -105.6544
23.2597 | -105.6612
23.2634 | -105.6676
23.2671 | -105.6749
23.2708 | -105.6813
23.2745 | -105.6857
23.2782 | -105.69
23.2818 | -105.6961
23.2855 | -105.7023
23.2892 | -105.7068
23.2929 | -105.7111
23.2966 | -105.7174
23.3003 | -105.7252
23.304 | -105.7326
23.3077 | -105.7392
23.3114 | -105.7446
23.315 | -105.7501
23.3187 | -105.7567
23.3224 | -105.7633
23.3261 | -105.769
23.3298 | -105.7757
23.3335 | -105.7811
23.3372 | -105.7867
23.3409 | -105.7933
23.3446 | -105.7983
23.3482 | -105.8041
23.3519 | -105.8107
23.3556 | -105.8192
23.3593 | -105.8303
23.363 | -105.8385
23.3667 | -105.8473
23.3704 | -105.8541
23.3741 | -105.8623
23.3777 | -105.8722
23.3814 | -105.88
23.3851 | -105.8864
23.3888 | -105.8923
23.3925 | -105.8979
23.3962 | -105.9035
23.3999 | -105.9092
23.4036 | -105.9156
23.4073 | -105.922
23.4109 | -105.9246
23.4146 | -105.9306
23.4183 | -105.9366
23.422 | -105.9408
23.4257 | -105.9444
23.4294 | -105.9472
23.4331 | -105.9508
23.4368 | -105.9552
23.4405 | -105.959
23.4441 | -105.9624
23.4478 | -105.9647
23.4515 | -105.9678
23.4552 | -105.9701
23.4589 | -105.9709
23.4626 | -105.9723
23.4663 | -105.9746
23.47 | -105.9742
23.4736 | -105.9737
23.4773 | -105.9737
23.481 | -105.974
23.4847 | -105.973
23.4884 | -105.9712
23.4921 | -105.9702
23.4958 | -105.9714
23.4995 | -105.9714
23.5032 | -105.9707
23.5068 | -105.9716
23.5105 | -105.9747
23.5142 | -105.9763
23.5179 | -105.9776
23.5216 | -105.9796
23.5253 | -105.9821
23.529 | -105.9848
23.5327 | -105.9871
23.5364 | -105.9908
23.54 | -105.9948
23.5437 | -105.998
23.5474 | -106.0006
23.5511 | -106.0042
23.5548 | -106.0097
23.5585 | -106.0146
23.5622 | -106.019
23.5659 | -106.0262
23.5696 | -106.0317
23.5732 | -106.0356
23.5769 | -106.0422
23.5806 | -106.0497
23.5843 | -106.057
23.588 | -106.0634
23.5917 | -106.0713
23.5954 | -106.0799
23.5991 | -106.0872
23.6027 | -106.0947
23.6064 | -106.1024
23.6101 | -106.1093
23.6138 | -106.1154
23.6175 | -106.1226
23.6212 | -106.1297
23.6249 | -106.1373
23.6286 | -106.1428
23.6323 | -106.1463
23.6359 | -106.1501
23.6396 | -106.1552
23.6433 | -106.1599
23.647 | -106.1645
23.6507 | -106.1687
23.6544 | -106.1717
23.6581 | -106.1729
23.6618 | -106.1739
23.6655 | -106.1764
23.6691 | -106.1777
23.6728 | -106.1799
23.6765 | -106.1812
23.6802 | -106.1825
23.6839 | -106.1829
23.6876 | -106.1817
23.6913 | -106.1799
23.695 | -106.1789
23.6986 | -106.1771
23.7023 | -106.1745
23.706 | -106.1733
23.7097 | -106.1717
23.7134 | -106.1716
23.7171 | -106.173
23.7208 | -106.1732
23.7245 | -106.1735
23.7282 | -106.1733
23.7318 | -106.1732
23.7355 | -106.1738
23.7392 | -106.1765
23.7429 | -106.1785
23.7466 | -106.1801
23.7503 | -106.1795
23.754 | -106.1801
23.7577 | -106.1841
23.7614 | -106.1859
23.765 | -106.1856
23.7687 | -106.1863
23.7724 | -106.1877
23.7761 | -106.1917
23.7798 | -106.1961
23.7835 | -106.1998
23.7872 | -106.2022
23.7909 | -106.2043
23.7946 | -106.2076
23.7982 | -106.2118
23.8019 | -106.2144
23.8056 | -106.2167
23.8093 | -106.2195
23.813 | -106.2221
23.8167 | -106.2251
23.8204 | -106.2285
23.8241 | -106.2321
23.8277 | -106.2365
23.8314 | -106.239
23.8351 | -106.2413
23.8388 | -106.243
23.8425 | -106.2433
23.8462 | -106.2444
23.8499 | -106.2475
23.8536 | -106.2512
23.8573 | -106.2555
23.8609 | -106.2589
23.8646 | -106.2619
23.8683 | -106.2662
23.872 | -106.2703
23.8757 | -106.2745
23.8794 | -106.2782
23.8831 | -106.2831
23.8868 | -106.2875
23.8905 | -106.2908
23.8941 | -106.2947
23.8978 | -106.2983
23.9015 | -106.2997
23.9052 | -106.3037
23.9089 | -106.3077
23.9126 | -106.3102
23.9163 | -106.312
23.92 | -106.3142
23.9236 | -106.3176
23.9273 | -106.3181
23.931 | -106.3202
23.9347 | -106.3244
23.9384 | -106.3285
23.9421 | -106.3308
23.9458 | -106.3289
23.9495 | -106.3302
23.9532 | -106.3315
23.9568 | -106.3328
23.9605 | -106.334
23.9642 | -106.3341
23.9679 | -106.3354
23.9716 | -106.3358
23.9753 | -106.335
23.979 | -106.3334
23.9827 | -106.3319
23.9864 | -106.3295
23.99 | -106.3275
23.9937 | -106.3266
23.9974 | -106.3252
24.0011 | -106.3228
24.0048 | -106.3201
24.0085 | -106.3175
24.0122 | -106.3139
24.0159 | -106.3113
24.0195 | -106.3088
24.0232 | -106.3069
24.0269 | -106.3038
24.0306 | -106.3001
24.0343 | -106.2979
24.038 | -106.2962
24.0417 | -106.2935
24.0454 | -106.2902
24.0491 | -106.288
24.0527 | -106.2851
24.0564 | -106.2818
24.0601 | -106.281
24.0638 | -106.2806
24.0675 | -106.281
24.0712 | -106.2801
24.0749 | -106.2803
24.0786 | -106.28
24.0823 | -106.2794
24.0859 | -106.28
24.0896 | -106.2825
24.0933 | -106.2845
24.097 | -106.2856
24.1007 | -106.2896
24.1044 | -106.2944
24.1081 | -106.2987
24.1118 | -106.304
24.1154 | -106.3063
24.1191 | -106.3104
24.1228 | -106.316
24.1265 | -106.3224
24.1302 | -106.3277
24.1339 | -106.3323
24.1376 | -106.3374
24.1413 | -106.3432
24.145 | -106.3501
24.1486 | -106.3562
24.1523 | -106.3623
24.156 | -106.3697
24.1597 | -106.3761
24.1634 | -106.3816
24.1671 | -106.3855
24.1708 | -106.3904
24.1745 | -106.3964
24.1782 | -106.4009
24.1818 | -106.4065
24.1855 | -106.4111
24.1892 | -106.4138
24.1929 | -106.4186
24.1966 | -106.4242
24.2003 | -106.43
24.204 | -106.4358
24.2077 | -106.4411
24.2114 | -106.4455
24.215 | -106.4491
24.2187 | -106.4523
24.2224 | -106.4568
24.2261 | -106.4621
24.2298 | -106.4671
24.2335 | -106.4708
24.2372 | -106.4731
24.2409 | -106.4766
24.2445 | -106.4794
24.2482 | -106.4829
24.2519 | -106.4857
24.2556 | -106.49
24.2593 | -106.4943
24.263 | -106.4983
24.2667 | -106.5039
24.2704 | -106.5071
24.2741 | -106.5084
24.2777 | -106.5119
24.2814 | -106.5157
24.2851 | -106.5174
24.2888 | -106.5196
24.2925 | -106.5225
24.2962 | -106.5184
24.2999 | -106.5072
24.3036 | -106.4928
24.3073 | -106.4815
24.3109 | -106.4691
24.3146 | -106.4558
24.3183 | -106.445
24.322 | -106.4359
24.3257 | -106.4264
24.3294 | -106.4151
24.3331 | -106.4018
24.3368 | -106.3884
24.3404 | -106.3757
24.3441 | -106.3611
24.3478 | -106.343
24.3515 | -106.3221
24.3552 | -106.3019
24.3589 | -106.2803
24.3626 | -106.2569
24.3663 | -106.2317
24.37 | -106.2063
24.3736 | -106.1808
24.3773 | -106.154
24.381 | -106.126
24.3847 | -106.0978
24.3884 | -106.0681
24.3921 | -106.0353
24.3958 | -106.0024
24.3995 | -105.9698
24.4032 | -105.9345
24.4068 | -105.9006
24.4105 | -105.868
24.4142 | -105.8366
24.4179 | -105.8043
24.4216 | -105.7715
24.4253 | -105.7412
24.429 | -105.7098
24.4327 | -105.6802
24.4364 | -105.6527
24.44 | -105.6241
24.4437 | -105.5942
24.4474 | -105.5687
24.4511 | -105.5456
24.4548 | -105.5225
24.4585 | -105.5029
24.4622 | -105.4854
24.4659 | -105.4684
24.4695 | -105.4527
24.4732 | -105.4418
24.4769 | -105.4487
24.4806 | -105.4425
24.4843 | -105.4367
24.488 | -105.4333
24.4917 | -105.4304
24.4954 | -105.4295
24.4991 | -105.4301
24.5027 | -105.4314
24.5064 | -105.4359
24.5101 | -105.444
24.5138 | -105.4534
24.5175 | -105.4625
24.5212 | -105.4736
24.5249 | -105.4872
24.5286 | -105.5007
24.5323 | -105.5146
24.5359 | -105.5312
24.5396 | -105.5474
24.5433 | -105.5634
24.547 | -105.5807
24.5507 | -105.5991
24.5544 | -105.6187
24.5581 | -105.6403
24.5618 | -105.6625
24.5654 | -105.6847
24.5691 | -105.7043
24.5728 | -105.7216
24.5765 | -105.736
24.5802 | -105.7453
24.5839 | -105.7513
24.5876 | -105.749
24.5913 | -105.7174
24.595 | -105.6604
24.5986 | -105.6046
24.6023 | -105.5506
24.606 | -105.5032
24.6097 | -105.4582
24.6134 | -105.4088
24.6171 | -105.3562
24.6208 | -105.2994
24.6245 | -105.2362
24.6282 | -105.1707
24.6318 | -105.1037
24.6355 | -105.0335
24.6392 | -104.9609
24.6429 | -104.8879
24.6466 | -104.8132
24.6503 | -104.7345
24.654 | -104.6484
24.6577 | -104.5583
24.6613 | -104.4651
24.665 | -104.3662
24.6687 | -104.2637
24.6724 | -104.1599
24.6761 | -104.0553
24.6798 | -103.95
24.6835 | -103.8444
24.6872 | -103.7385
24.6909 | -103.6339
24.6945 | -103.5325
24.6982 | -103.4331
24.7019 | -103.3342
24.7056 | -103.2376
24.7093 | -103.1445
24.713 | -103.0549
24.7167 | -102.9683
24.7204 | -102.8857
24.7241 | -102.8071
24.7277 | -102.732
24.7314 | -102.6616
24.7351 | -102.5965
24.7388 | -102.5378
24.7425 | -102.4844
24.7462 | -102.4359
24.7499 | -102.3938
24.7536 | -102.3563
24.7572 | -102.3242
24.7609 | -102.2985
24.7646 | -102.279
24.7683 | -102.2658
24.772 | -102.2586
24.7757 | -102.258
24.7794 | -102.2641
24.7831 | -102.2756
24.7868 | -102.2937
24.7904 | -102.3186
24.7941 | -102.3489
24.7978 | -102.3848
24.8015 | -102.4271
24.8052 | -102.4752
24.8089 | -102.529
24.8126 | -102.5884
24.8163 | -102.6528
24.82 | -102.7223
24.8236 | -102.7971
24.8273 | -102.876
24.831 | -102.9595
24.8347 | -103.0477
24.8384 | -103.1396
24.8421 | -103.2352
24.8458 | -103.3344
24.8495 | -103.4362
24.8532 | -103.5408
24.8568 | -103.6467
24.8605 | -103.754
24.8642 | -103.8631
24.8679 | -103.974
24.8716 | -104.0847
24.8753 | -104.1954
24.879 | -104.3065
24.8827 | -104.4174
24.8863 | -104.5254
24.89 | -104.632
24.8937 | -104.7376
24.8974 | -104.841
24.9011 | -104.9422
24.9048 | -105.0398
24.9085 | -105.1337
24.9122 | -105.2249
24.9159 | -105.313
24.9195 | -105.4002
24.9232 | -105.4861
24.9269 | -105.5696
24.9306 | -105.6527
24.9343 | -105.7325
24.938 | -105.8124
24.9417 | -105.9105
24.9454 | -106.0214
24.9491 | -106.1149
24.9527 | -106.1903
24.9564 | -106.2442
24.9601 | -106.2848
24.9638 | -106.3215
24.9675 | -106.3525
24.9712 | -106.3789
24.9749 | -106.4038
24.9786 | -106.4235
24.9822 | -106.4401
24.9859 | -106.4517
24.9896 | -106.4583
24.9933 | -106.4636
24.997 | -106.4682
25.0007 | -106.4735
25.0044 | -106.4785
25.0081 | -106.481
25.0118 | -106.4871
25.0154 | -106.491
25.0191 | -106.4957
25.0228 | -106.5002
25.0265 | -106.5034
25.0302 | -106.5076
25.0339 | -106.5124
25.0376 | -106.5154
25.0413 | -106.5177
25.045 | -106.519
25.0486 | -106.5191
25.0523 | -106.5196
25.056 | -106.5215
25.0597 | -106.5238
25.0634 | -106.5242
25.0671 | -106.5239
25.0708 | -106.524
25.0745 | -106.5226
25.0782 | -106.5214
25.0818 | -106.5225
25.0855 | -106.5239
25.0892 | -106.524
25.0929 | -106.5228
25.0966 | -106.5213
25.1003 | -106.5204
25.104 | -106.5204
25.1077 | -106.5229
25.1113 | -106.5244
25.115 | -106.526
25.1187 | -106.5278
25.1224 | -106.5306
25.1261 | -106.5313
25.1298 | -106.5319
25.1335 | -106.5315
25.1372 | -106.5335
25.1409 | -106.536
25.1445 | -106.539
25.1482 | -106.543
25.1519 | -106.5469
25.1556 | -106.5496
25.1593 | -106.5524
25.163 | -106.5562
25.1667 | -106.5612
25.1704 | -106.5673
25.1741 | -106.5711
25.1777 | -106.5748
25.1814 | -106.5797
25.1851 | -106.5834
25.1888 | -106.5878
25.1925 | -106.5925
25.1962 | -106.5981
25.1999 | -106.6026
25.2036 | -106.6086
25.2072 | -106.6146
25.2109 | -106.6199
25.2146 | -106.625
25.2183 | -106.6304
25.222 | -106.6346
25.2257 | -106.6393
25.2294 | -106.6437
25.2331 | -106.6469
25.2368 | -106.65
25.2404 | -106.6545
25.2441 | -106.659
25.2478 | -106.6623
25.2515 | -106.6643
25.2552 | -106.6663
25.2589 | -106.6695
25.2626 | -106.671
25.2663 | -106.6706
25.27 | -106.6712
25.2736 | -106.6711
25.2773 | -106.6688
25.281 | -106.6661
25.2847 | -106.6619
25.2884 | -106.659
25.2921 | -106.6577
25.2958 | -106.6555
25.2995 | -106.652
25.3031 | -106.6473
25.3068 | -106.6441
25.3105 | -106.6408
25.3142 | -106.6365
25.3179 | -106.6326
25.3216 | -106.6302
25.3253 | -106.6272
25.329 | -106.6234
25.3327 | -106.618
25.3363 | -106.6116
25.34 | -106.606
25.3437 | -106.5998
25.3474 | -106.5957
25.3511 | -106.5928
25.3548 | -106.5875
25.3585 | -106.5806
25.3622 | -106.5756
25.3659 | -106.5701
25.3695 | -106.5649
25.3732 | -106.5593
25.3769 | -106.5523
25.3806 | -106.547
25.3843 | -106.5432
25.388 | -106.5373
25.3917 | -106.5323
25.3954 | -106.5286
25.399 | -106.5247
25.4027 | -106.5211
25.4064 | -106.5174
25.4101 | -106.5137
25.4138 | -106.5098
25.4175 | -106.5052
25.4212 | -106.5013
25.4249 | -106.4986
25.4286 | -106.496
25.4322 | -106.4926
25.4359 | -106.4888
25.4396 | -106.4853
25.4433 | -106.4813
25.447 | -106.4769
25.4507 | -106.4711
25.4544 | -106.4659
25.4581 | -106.4604
25.4618 | -106.4543
25.4654 | -106.4471
25.4691 | -106.4413
25.4728 | -106.435
25.4765 | -106.4298
25.4802 | -106.4242
25.4839 | -106.4202
25.4876 | -106.4169
25.4913 | -106.4139
25.495 | -106.4102
25.4986 | -106.4062
25.5023 | -106.4012
25.506 | -106.3961
25.5097 | -106.391
25.5134 | -106.3868
25.5171 | -106.3831
25.5208 | -106.3792
25.5245 | -106.3762
25.5281 | -106.3734
25.5318 | -106.373
25.5355 | -106.373
25.5392 | -106.3768
25.5429 | -106.3765
25.5466 | -106.3771
25.5503 | -106.3764
25.554 | -106.3771
25.5577 | -106.3782
25.5613 | -106.3799
25.565 | -106.3815
25.5687 | -106.3838
25.5724 | -106.3882
25.5761 | -106.393
25.5798 | -106.3968
25.5835 | -106.4009
25.5872 | -106.4031
25.5909 | -106.407
25.5945 | -106.4123
25.5982 | -106.4176
25.6019 | -106.4237
25.6056 | -106.4297
25.6093 | -106.4356
25.613 | -106.4423
25.6167 | -106.4484
25.6204 | -106.454
25.624 | -106.4622
25.6277 | -106.4701
25.6314 | -106.4765
25.6351 | -106.4811
25.6388 | -106.4881
25.6425 | -106.4956
25.6462 | -106.5037
25.6499 | -106.5114
25.6536 | -106.5174
25.6572 | -106.5224
25.6609 | -106.5274
25.6646 | -106.5331
25.6683 | -106.5385
25.672 | -106.5431
25.6757 | -106.5467
25.6794 | -106.5501
25.6831 | -106.554
25.6868 | -106.5579
25.6904 | -106.5617
25.6941 | -106.5674
25.6978 | -106.571
25.7015 | -106.5737
25.7052 | -106.5785
25.7089 | -106.5826
25.7126 | -106.585
25.7163 | -106.5887
25.72 | -106.593
25.7236 | -106.5946
25.7273 | -106.5961
25.731 | -106.5963
25.7347 | -106.5993
25.7384 | -106.6042
25.7421 | -106.606
25.7458 | -106.6066
25.7495 | -106.6068
25.7531 | -106.6074
25.7568 | -106.6088
25.7605 | -106.6093
25.7642 | -106.6085
25.7679 | -106.6084
25.7716 | -106.6088
25.7753 | -106.609
25.779 | -106.6103
25.7827 | -106.6108
25.7863 | -106.6097
25.79 | -106.6104
25.7937 | -106.6115
25.7974 | -106.6121
25.8011 | -106.6115
25.8048 | -106.6119
25.8085 | -106.6116
25.8122 | -106.6118
25.8159 | -106.6136
25.8195 | -106.6144
25.8232 | -106.6141
25.8269 | -106.6134
25.8306 | -106.6118
25.8343 | -106.6102
25.838 | -106.6076
25.8417 | -106.6049
25.8454 | -106.6021
25.849 | -106.5983
25.8527 | -106.5946
25.8564 | -106.5909
25.8601 | -106.588
25.8638 | -106.5859
25.8675 | -106.5829
25.8712 | -106.5791
25.8749 | -106.5759
25.8786 | -106.5721
25.8822 | -106.5669
25.8859 | -106.5624
25.8896 | -106.5579
25.8933 | -106.5526
25.897 | -106.5469
25.9007 | -106.5403
25.9044 | -106.5334
25.9081 | -106.528
25.9118 | -106.5234
25.9154 | -106.519
25.9191 | -106.5128
25.9228 | -106.5067
25.9265 | -106.4997
25.9302 | -106.4923
25.9339 | -106.4851
25.9376 | -106.4778
25.9413 | -106.4719
25.9449 | -106.4676
25.9486 | -106.4627
25.9523 | -106.458
25.956 | -106.4527
25.9597 | -106.4454
25.9634 | -106.4395
25.9671 | -106.4338
25.9708 | -106.4286
25.9745 | -106.4248
25.9781 | -106.4209
25.9818 | -106.4151
25.9855 | -106.4117
25.9892 | -106.4086
25.9929 | -106.4072
25.9966 | -106.406
26.0003 | -106.4063
26.004 | -106.406
26.0077 | -106.405
26.0113 | -106.403
26.015 | -106.4012
26.0187 | -106.3996
26.0224 | -106.3994
26.0261 | -106.3999
26.0298 | -106.4022
26.0335 | -106.4045
26.0372 | -106.4074
26.0408 | -106.4106
26.0445 | -106.4135
26.0482 | -106.4161
26.0519 | -106.4202
26.0556 | -106.4238
26.0593 | -106.4277
26.063 | -106.4335
26.0667 | -106.4402
26.0704 | -106.445
26.074 | -106.4538
26.0777 | -106.4584
26.0814 | -106.4641
26.0851 | -106.4694
26.0888 | -106.475
26.0925 | -106.4807
26.0962 | -106.4863
26.0999 | -106.4944
26.1036 | -106.5004
26.1072 | -106.5037
26.1109 | -106.508
26.1146 | -106.5136
26.1183 | -106.5198
26.122 | -106.5254
26.1257 | -106.5297
26.1294 | -106.5347
26.1331 | -106.5402
26.1368 | -106.546
26.1404 | -106.5499
26.1441 | -106.5546
26.1478 | -106.558
26.1515 | -106.5609
26.1552 | -106.5636
26.1589 | -106.5654
26.1626 | -106.5698
26.1663 | -106.5737
26.1699 | -106.5766
26.1736 | -106.5783
26.1773 | -106.5803
26.181 | -106.5808
26.1847 | -106.5815
26.1884 | -106.5839
26.1921 | -106.5851
26.1958 | -106.5875
26.1995 | -106.589
26.2031 | -106.5905
26.2068 | -106.5903
26.2105 | -106.5909
26.2142 | -106.5918
26.2179 | -106.5915
26.2216 | -106.5885
26.2253 | -106.5865
26.229 | -106.584
26.2327 | -106.5806
26.2363 | -106.5785
26.24 | -106.5762
26.2437 | -106.5725
26.2474 | -106.5703
26.2511 | -106.5682
26.2548 | -106.5645
26.2585 | -106.561
26.2622 | -106.5573
26.2658 | -106.5542
26.2695 | -106.5516
26.2732 | -106.5488
26.2769 | -106.5441
26.2806 | -106.5384
26.2843 | -106.5325
26.288 | -106.5286
26.2917 | -106.5248
26.2954 | -106.5217
26.299 | -106.5192
26.3027 | -106.5158
26.3064 | -106.5126
26.3101 | -106.5085
26.3138 | -106.5033
26.3175 | -106.4996
26.3212 | -106.498
26.3249 | -106.4948
26.3286 | -106.4936
26.3322 | -106.4937
26.3359 | -106.4927
26.3396 | -106.491
26.3433 | -106.4895
26.347 | -106.4884
26.3507 | -106.4893
26.3544 | -106.487
26.3581 | -106.4859
26.3618 | -106.4855
26.3654 | -106.4861
26.3691 | -106.4848
26.3728 | -106.4833
26.3765 | -106.4856
26.3802 | -106.4901
26.3839 | -106.4924
26.3876 | -106.493
26.3913 | -106.4948
26.3949 | -106.4971
26.3986 | -106.499
26.4023 | -106.5006
26.406 | -106.5029
26.4097 | -106.505
26.4134 | -106.5063
26.4171 | -106.5076
26.4208 | -106.5093
26.4245 | -106.5125
26.4281 | -106.5145
26.4318 | -106.5173
26.4355 | -106.5188
26.4392 | -106.5208
26.4429 | -106.5203
26.4466 | -106.5194
26.4503 | -106.5214
26.454 | -106.5245
26.4577 | -106.5271
26.4613 | -106.5295
26.465 | -106.5324
26.4687 | -106.533
26.4724 | -106.5326
26.4761 | -106.5334
26.4798 | -106.5362
26.4835 | -106.5375
26.4872 | -106.5373
26.4908 | -106.5368
26.4945 | -106.5354
26.4982 | -106.5356
26.5019 | -106.5354
26.5056 | -106.5354
26.5093 | -106.5343
26.513 | -106.5329
26.5167 | -106.5306
26.5204 | -106.5299
26.524 | -106.5296
26.5277 | -106.5277
26.5314 | -106.5257
26.5351 | -106.5218
26.5388 | -106.5185
26.5425 | -106.5154
26.5462 | -106.5127
26.5499 | -106.5093
26.5536 | -106.5091
26.5572 | -106.5054
26.5609 | -106.4998
26.5646 | -106.4957
26.5683 | -106.4907
26.572 | -106.485
26.5757 | -106.4811
26.5794 | -106.4788
26.5831 | -106.4757
26.5867 | -106.4705
26.5904 | -106.4668
26.5941 | -106.4654
26.5978 | -106.4634
26.6015 | -106.4607
26.6052 | -106.4594
26.6089 | -106.4577
26.6126 | -106.455
26.6163 | -106.4548
26.6199 | -106.4548
26.6236 | -106.4514
26.6273 | -106.4485
26.631 | -106.4476
26.6347 | -106.4479
26.6384 | -106.4481
26.6421 | -106.448
26.6458 | -106.4486
26.6495 | -106.4498
26.6531 | -106.4512
26.6568 | -106.4538
26.6605 | -106.4569
26.6642 | -106.459
26.6679 | -106.4616
26.6716 | -106.4648
26.6753 | -106.4684
26.679 | -106.4727
26.6826 | -106.4767
26.6863 | -106.4804
26.69 | -106.4859
26.6937 | -106.4911
26.6974 | -106.4962
26.7011 | -106.501
26.7048 | -106.5065
26.7085 | -106.5108
26.7122 | -106.5156
26.7158 | -106.5227
26.7195 | -106.5299
26.7232 | -106.5371
26.7269 | -106.5431
26.7306 | -106.5503
26.7343 | -106.5569
26.738 | -106.5634
26.7417 | -106.5701
26.7454 | -106.5785
26.749 | -106.587
26.7527 | -106.5935
26.7564 | -106.5987
26.7601 | -106.6068
26.7638 | -106.6159
26.7675 | -106.623
26.7712 | -106.6287
26.7749 | -106.6339
26.7786 | -106.6392
26.7822 | -106.6446
26.7859 | -106.65
26.7896 | -106.6565
26.7933 | -106.6617
26.797 | -106.6663
26.8007 | -106.6706
26.8044 | -106.6757
26.8081 | -106.6801
26.8117 | -106.6851
26.8154 | -106.6917
26.8191 | -106.6986
26.8228 | -106.7052
26.8265 | -106.7084
26.8302 | -106.7107
26.8339 | -106.714
26.8376 | -106.7186
26.8413 | -106.7223
26.8449 | -106.7259
26.8486 | -106.7279
26.8523 | -106.7308
26.856 | -106.7333
26.8597 | -106.7349
26.8634 | -106.7356
26.8671 | -106.7389
26.8708 | -106.7416
26.8745 | -106.7412
26.8781 | -106.7425
26.8818 | -106.7437
26.8855 | -106.7438
26.8892 | -106.7439
26.8929 | -106.744
26.8966 | -106.7429
26.9003 | -106.7421
26.904 | -106.7409
26.9076 | -106.7398
26.9113 | -106.7395
26.915 | -106.7387
26.9187 | -106.7373
26.9224 | -106.7368
26.9261 | -106.7366
26.9298 | -106.7346
26.9335 | -106.7328
26.9372 | -106.732
26.9408 | -106.7305
26.9445 | -106.7301
26.9482 | -106.73
26.9519 | -106.7283
26.9556 | -106.7269
26.9593 | -106.7254
26.963 | -106.725
26.9667 | -106.7248
26.9704 | -106.7238
26.974 | -106.7214
26.9777 | -106.72
26.9814 | -106.7186
26.9851 | -106.7169
26.9888 | -106.7157
26.9925 | -106.7148
26.9962 | -106.7124
26.9999 | -106.7103
27.0036 | -106.7101
27.0072 | -106.7112
27.0109 | -106.7114
27.0146 | -106.7105
27.0183 | -106.7096
27.022 | -106.7077
27.0257 | -106.7073
27.0294 | -106.7059
27.0331 | -106.7048
27.0367 | -106.7049
27.0404 | -106.7049
27.0441 | -106.7046
27.0478 | -106.7038
27.0515 | -106.7039
27.0552 | -106.7036
27.0589 | -106.7016
27.0626 | -106.6987
27.0663 | -106.6973
27.0699 | -106.6963
27.0736 | -106.697
27.0773 | -106.6971
27.081 | -106.6965
27.0847 | -106.6954
27.0884 | -106.6943
27.0921 | -106.6907
27.0958 | -106.6864
27.0995 | -106.6847
27.1031 | -106.6835
27.1068 | -106.6834
27.1105 | -106.6828
27.1142 | -106.6805
27.1179 | -106.6781
27.1216 | -106.6759
27.1253 | -106.6724
27.129 | -106.6682
27.1326 | -106.6651
27.1363 | -106.6593
27.14 | -106.6564
27.1437 | -106.6535
27.1474 | -106.6511
27.1511 | -106.6483
27.1548 | -106.646
27.1585 | -106.6442
27.1622 | -106.6415
27.1658 | -106.6383
27.1695 | -106.6365
27.1732 | -106.6329
27.1769 | -106.6296
27.1806 | -106.6266
27.1843 | -106.6224
27.188 | -106.6188
27.1917 | -106.6157
27.1954 | -106.6127
27.199 | -106.6088
27.2027 | -106.6056
27.2064 | -106.602
27.2101 | -106.6001
27.2138 | -106.5969
27.2175 | -106.5939
27.2212 | -106.5908
27.2249 | -106.5885
27.2285 | -106.587
27.2322 | -106.5853
27.2359 | -106.5836
27.2396 | -106.5796
27.2433 | -106.5768
27.247 | -106.5754
27.2507 | -106.5737
27.2544 | -106.5726
27.2581 | -106.5731
27.2617 | -106.5727
27.2654 | -106.5713
27.2691 | -106.5705
27.2728 | -106.5712
27.2765 | -106.5715
27.2802 | -106.5713
27.2839 | -106.5709
27.2876 | -106.5725
27.2913 | -106.5728
27.2949 | -106.5723
27.2986 | -106.5707
27.3023 | -106.5698
27.306 | -106.5679
27.3097 | -106.5669
27.3134 | -106.5659
27.3171 | -106.5658
27.3208 | -106.5659
27.3244 | -106.5655
27.3281 | -106.5636
27.3318 | -106.5619
27.3355 | -106.5606
27.3392 | -106.5586
27.3429 | -106.5559
27.3466 | -106.553
27.3503 | -106.5485
27.354 | -106.545
27.3576 | -106.542
27.3613 | -106.5398
27.365 | -106.5384
27.3687 | -106.5362
27.3724 | -106.5319
27.3761 | -106.5253
27.3798 | -106.5189
27.3835 | -106.5138
27.3872 | -106.5095
27.3908 | -106.5032
27.3945 | -106.498
27.3982 | -106.4928
27.4019 | -106.4877
27.4056 | -106.4828
27.4093 | -106.4753
27.413 | -106.4693
27.4167 | -106.4641
27.4204 | -106.4578
27.424 | -106.453
27.4277 | -106.4472
27.4314 | -106.4418
27.4351 | -106.4373
27.4388 | -106.4331
27.4425 | -106.4302
27.4462 | -106.4285
27.4499 | -106.4265
27.4535 | -106.4231
27.4572 | -106.4202
27.4609 | -106.4177
27.4646 | -106.4165
27.4683 | -106.4158
27.472 | -106.4147
27.4757 | -106.4155
27.4794 | -106.417
27.4831 | -106.4161
27.4867 | -106.415
27.4904 | -106.4154
27.4941 | -106.4167
27.4978 | -106.4191
27.5015 | -106.4204
27.5052 | -106.42
27.5089 | -106.422
27.5126 | -106.4256
27.5163 | -106.4286
27.5199 | -106.4322
27.5236 | -106.4377
27.5273 | -106.4426
27.531 | -106.446
27.5347 | -106.4499
27.5384 | -106.455
27.5421 | -106.4601
27.5458 | -106.4645
27.5494 | -106.4699
27.5531 | -106.4752
27.5568 | -106.4787
27.5605 | -106.4839
27.5642 | -106.4898
27.5679 | -106.4946
27.5716 | -106.4997
27.5753 | -106.5048
27.579 | -106.5105
27.5826 | -106.5161
27.5863 | -106.5214
27.59 | -106.5274
27.5937 | -106.534
27.5974 | -106.5391
27.6011 | -106.546
27.6048 | -106.5513
27.6085 | -106.5558
27.6122 | -106.5607
27.6158 | -106.5657
27.6195 | -106.5721
27.6232 | -106.5772
27.6269 | -106.5843
27.6306 | -106.5912
27.6343 | -106.5959
27.638 | -106.5985
27.6417 | -106.6001
27.6454 | -106.6022
27.649 | -106.6049
27.6527 | -106.6063
27.6564 | -106.6083
27.6601 | -106.6097
27.6638 | -106.6095
27.6675 | -106.6075
27.6712 | -106.6087
27.6749 | -106.6085
27.6785 | -106.6065
27.6822 | -106.6056
27.6859 | -106.6054
27.6896 | -106.6049
27.6933 | -106.6038
27.697 | -106.602
27.7007 | -106.5978
27.7044 | -106.5936
27.7081 | -106.5901
27.7117 | -106.5866
27.7154 | -106.5832
27.7191 | -106.5796
27.7228 | -106.5747
27.7265 | -106.569
27.7302 | -106.5641
27.7339 | -106.5594
27.7376 | -106.5556
27.7413 | -106.5527
27.7449 | -106.5465
27.7486 | -106.5406
27.7523 | -106.5337
27.756 | -106.5269
27.7597 | -106.5192
27.7634 | -106.5129
27.7671 | -106.5075
27.7708 | -106.5033
27.7744 | -106.4984
27.7781 | -106.492
27.7818 | -106.4857
27.7855 | -106.4768
27.7892 | -106.4674
27.7929 | -106.4597
27.7966 | -106.4536
27.8003 | -106.4492
27.804 | -106.4446
27.8076 | -106.4389
27.8113 | -106.4326
27.815 | -106.4245
27.8187 | -106.418
27.8224 | -106.4126
27.8261 | -106.4072
27.8298 | -106.4023
27.8335 | -106.3965
27.8372 | -106.3922
27.8408 | -106.388
27.8445 | -106.3849
27.8482 | -106.3811
27.8519 | -106.3778
27.8556 | -106.3741
27.8593 | -106.3695
27.863 | -106.3655
27.8667 | -106.3634
27.8703 | -106.3617
27.874 | -106.3618
27.8777 | -106.3634
27.8814 | -106.3636
27.8851 | -106.3633
27.8888 | -106.3642
27.8925 | -106.3663
27.8962 | -106.3692
27.8999 | -106.3713
27.9035 | -106.3726
27.9072 | -106.3747
27.9109 | -106.3766
27.9146 | -106.3796
27.9183 | -106.3824
27.922 | -106.3848
27.9257 | -106.3854
27.9294 | -106.3846
27.9331 | -106.3857
27.9367 | -106.3881
27.9404 | -106.3927
27.9441 | -106.3965
27.9478 | -106.4003
27.9515 | -106.4041
27.9552 | -106.4068
27.9589 | -106.4091
27.9626 | -106.4119
27.9662 | -106.4147
27.9699 | -106.4177
27.9736 | -106.4204
27.9773 | -106.4219
27.981 | -106.4225
27.9847 | -106.4234
27.9884 | -106.4244
27.9921 | -106.4222
27.9958 | -106.4213
27.9994 | -106.4215
28.0031 | -106.4208