## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 209 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 206
Hood | 02-92B/P | 146
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1014.8 | hPa |  | PRC PRO 01
Humidity | 46.4 | %/Hr
Internal temperature of calibration bench | 23.51 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1021 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 25.8 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.95 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.05 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3139 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 19/04/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 20/04/2018
Engineer | Anthony HUE | Approval |  | 23/04/2018
CEO | Robert GLEMOT | Approval |  | 23/04/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5879 | 57.4782
0.012209 | -36.1642 | 38.3067
0.018313 | -35.105 | 28.1776
0.024417 | -34.6412 | 21.9434
0.030521 | -34.5287 | 18.1852
0.036626 | -34.3835 | 15.1353
0.04273 | -34.2873 | 12.7749
0.048834 | -34.2412 | 11.2586
0.054938 | -34.1768 | 10.1244
0.061043 | -34.116 | 9.0458
0.067147 | -34.1056 | 8.227
0.073251 | -34.1062 | 7.4708
0.079355 | -34.124 | 6.8122
0.08546 | -34.1373 | 6.2203
0.091564 | -34.1321 | 5.7815
0.097668 | -34.1139 | 5.5338
0.10377 | -34.0917 | 5.2866
0.10988 | -34.0743 | 5.0095
0.11598 | -34.065 | 4.7166
0.12209 | -34.0668 | 4.4321
0.12819 | -34.0766 | 4.1481
0.13429 | -34.0874 | 3.8602
0.1404 | -34.0908 | 3.5827
0.1465 | -34.0929 | 3.3256
0.15261 | -34.0904 | 3.1059
0.15871 | -34.0841 | 2.9436
0.16482 | -34.0761 | 2.84
0.17092 | -34.0737 | 2.7996
0.17702 | -34.0714 | 2.7266
0.18313 | -34.0703 | 2.6481
0.18923 | -34.0697 | 2.5396
0.19534 | -34.0682 | 2.4135
0.20144 | -34.0667 | 2.2725
0.20754 | -34.0638 | 2.1527
0.21365 | -34.0626 | 2.039
0.21975 | -34.0621 | 1.9418
0.22586 | -34.063 | 1.8549
0.23196 | -34.0653 | 1.7827
0.23807 | -34.0693 | 1.72
0.24417 | -34.0732 | 1.6613
0.25027 | -34.0771 | 1.5991
0.25638 | -34.0794 | 1.5372
0.26248 | -34.0811 | 1.4708
0.26859 | -34.0819 | 1.4057
0.27469 | -34.0822 | 1.335
0.2808 | -34.0832 | 1.2619
0.2869 | -34.0839 | 1.1913
0.293 | -34.0848 | 1.1246
0.29911 | -34.0857 | 1.0647
0.30521 | -34.0866 | 1.0057
0.31132 | -34.0873 | 0.96663
0.31742 | -34.0885 | 0.92349
0.32353 | -34.0888 | 0.88818
0.32963 | -34.0896 | 0.8587
0.33573 | -34.0906 | 0.83798
0.34184 | -34.0919 | 0.81622
0.34794 | -34.094 | 0.79248
0.35405 | -34.096 | 0.76828
0.36015 | -34.0985 | 0.74293
0.36626 | -34.1004 | 0.71549
0.37236 | -34.1025 | 0.67892
0.37846 | -34.1046 | 0.63485
0.38457 | -34.1054 | 0.57776
0.39067 | -34.1064 | 0.53113
0.39678 | -34.107 | 0.47924
0.40288 | -34.1074 | 0.42517
0.40899 | -34.1072 | 0.37316
0.41509 | -34.1065 | 0.32215
0.42119 | -34.1064 | 0.27876
0.4273 | -34.1063 | 0.23249
0.4334 | -34.1061 | 0.18634
0.43951 | -34.1063 | 0.14302
0.44561 | -34.1071 | 0.10333
0.45172 | -34.1086 | 0.071778
0.45782 | -34.1093 | 0.034839
0.46392 | -34.1105 | -0.0012532
0.47003 | -34.1115 | -0.033507
0.47613 | -34.1123 | -0.067482
0.48224 | -34.1127 | -0.10216
0.48834 | -34.1118 | -0.13686
0.49445 | -34.1115 | -0.17208
0.50055 | -34.1109 | -0.20414
0.50665 | -34.1102 | -0.23834
0.51276 | -34.1091 | -0.27241
0.51886 | -34.1081 | -0.30642
0.52497 | -34.1068 | -0.33083
0.53107 | -34.106 | -0.36044
0.53717 | -34.1053 | -0.38881
0.54328 | -34.1044 | -0.41187
0.54938 | -34.1042 | -0.43457
0.55549 | -34.1042 | -0.45574
0.56159 | -34.1053 | -0.46772
0.5677 | -34.1056 | -0.48264
0.5738 | -34.1066 | -0.49927
0.5799 | -34.1074 | -0.51455
0.58601 | -34.1085 | -0.5251
0.59211 | -34.1101 | -0.53186
0.59822 | -34.1115 | -0.54176
0.60432 | -34.113 | -0.55424
0.61043 | -34.1142 | -0.56262
0.61653 | -34.1155 | -0.57332
0.62263 | -34.117 | -0.58312
0.62874 | -34.1177 | -0.59417
0.63484 | -34.1185 | -0.60385
0.64095 | -34.119 | -0.61446
0.64705 | -34.1194 | -0.62462
0.65316 | -34.1196 | -0.63741
0.65926 | -34.1197 | -0.65349
0.66536 | -34.1186 | -0.67106
0.67147 | -34.1183 | -0.68456
0.67757 | -34.1179 | -0.6991
0.68368 | -34.1171 | -0.71818
0.68978 | -34.1164 | -0.7356
0.69589 | -34.1154 | -0.75477
0.70199 | -34.1145 | -0.77762
0.70809 | -34.1136 | -0.797
0.7142 | -34.1129 | -0.81764
0.7203 | -34.1118 | -0.83703
0.72641 | -34.1111 | -0.85967
0.73251 | -34.1104 | -0.87986
0.73862 | -34.1106 | -0.90188
0.74472 | -34.1103 | -0.92212
0.75082 | -34.11 | -0.94442
0.75693 | -34.1103 | -0.96647
0.76303 | -34.1108 | -0.9887
0.76914 | -34.1126 | -1.0094
0.77524 | -34.1135 | -1.0293
0.78135 | -34.1147 | -1.0497
0.78745 | -34.1161 | -1.0699
0.79355 | -34.1173 | -1.0899
0.79966 | -34.119 | -1.1096
0.80576 | -34.1212 | -1.1295
0.81187 | -34.1228 | -1.1513
0.81797 | -34.1247 | -1.1727
0.82408 | -34.1264 | -1.1934
0.83018 | -34.1281 | -1.2141
0.83628 | -34.13 | -1.2328
0.84239 | -34.1311 | -1.2534
0.84849 | -34.1326 | -1.2752
0.8546 | -34.134 | -1.2952
0.8607 | -34.1352 | -1.316
0.86681 | -34.1363 | -1.3351
0.87291 | -34.1373 | -1.354
0.87901 | -34.1377 | -1.3721
0.88512 | -34.1385 | -1.3904
0.89122 | -34.1392 | -1.4087
0.89733 | -34.1395 | -1.4245
0.90343 | -34.14 | -1.4422
0.90953 | -34.1401 | -1.4567
0.91564 | -34.1401 | -1.4726
0.92174 | -34.1403 | -1.4877
0.92785 | -34.1404 | -1.5024
0.93395 | -34.1404 | -1.5164
0.94006 | -34.1406 | -1.5294
0.94616 | -34.1403 | -1.5414
0.95226 | -34.1404 | -1.5542
0.95837 | -34.1404 | -1.5682
0.96447 | -34.1406 | -1.5811
0.97058 | -34.1404 | -1.5946
0.97668 | -34.1402 | -1.6081
0.98279 | -34.1399 | -1.621
0.98889 | -34.1397 | -1.6318
0.99499 | -34.1396 | -1.6453
1.0011 | -34.1395 | -1.6561
1.0072 | -34.1391 | -1.6679
1.0133 | -34.1388 | -1.6788
1.0194 | -34.1383 | -1.6919
1.0255 | -34.138 | -1.7057
1.0316 | -34.1374 | -1.7178
1.0377 | -34.137 | -1.7316
1.0438 | -34.1364 | -1.7444
1.0499 | -34.1357 | -1.7598
1.056 | -34.135 | -1.7742
1.0621 | -34.1344 | -1.7886
1.0682 | -34.1339 | -1.8011
1.0743 | -34.1332 | -1.8167
1.0805 | -34.1327 | -1.8312
1.0866 | -34.1318 | -1.8463
1.0927 | -34.131 | -1.8609
1.0988 | -34.1303 | -1.8759
1.1049 | -34.1295 | -1.8902
1.111 | -34.1287 | -1.9056
1.1171 | -34.128 | -1.9203
1.1232 | -34.1276 | -1.9355
1.1293 | -34.1272 | -1.9496
1.1354 | -34.1269 | -1.9647
1.1415 | -34.1265 | -1.9803
1.1476 | -34.1259 | -1.9959
1.1537 | -34.1254 | -2.0106
1.1598 | -34.1255 | -2.0254
1.1659 | -34.1252 | -2.0388
1.172 | -34.125 | -2.0527
1.1781 | -34.1247 | -2.0668
1.1842 | -34.1246 | -2.081
1.1903 | -34.125 | -2.0921
1.1964 | -34.1249 | -2.1056
1.2025 | -34.125 | -2.1196
1.2086 | -34.1253 | -2.132
1.2147 | -34.1253 | -2.1446
1.2209 | -34.1257 | -2.1558
1.227 | -34.1264 | -2.1667
1.2331 | -34.1268 | -2.1787
1.2392 | -34.1272 | -2.1907
1.2453 | -34.1276 | -2.2018
1.2514 | -34.128 | -2.2118
1.2575 | -34.1286 | -2.2229
1.2636 | -34.1295 | -2.2321
1.2697 | -34.1301 | -2.2414
1.2758 | -34.1307 | -2.2527
1.2819 | -34.1314 | -2.2627
1.288 | -34.1321 | -2.2736
1.2941 | -34.1329 | -2.2823
1.3002 | -34.1336 | -2.2905
1.3063 | -34.1341 | -2.2997
1.3124 | -34.1348 | -2.3089
1.3185 | -34.1355 | -2.3181
1.3246 | -34.1362 | -2.3274
1.3307 | -34.137 | -2.3361
1.3368 | -34.1378 | -2.345
1.3429 | -34.1384 | -2.3542
1.349 | -34.1389 | -2.3636
1.3551 | -34.1396 | -2.3732
1.3613 | -34.1402 | -2.3816
1.3674 | -34.1408 | -2.3918
1.3735 | -34.1412 | -2.4017
1.3796 | -34.1416 | -2.411
1.3857 | -34.142 | -2.4204
1.3918 | -34.1425 | -2.4307
1.3979 | -34.143 | -2.44
1.404 | -34.1435 | -2.4514
1.4101 | -34.1439 | -2.4619
1.4162 | -34.1443 | -2.4726
1.4223 | -34.1448 | -2.4841
1.4284 | -34.1452 | -2.4942
1.4345 | -34.1455 | -2.5058
1.4406 | -34.1457 | -2.5193
1.4467 | -34.1462 | -2.531
1.4528 | -34.1466 | -2.5424
1.4589 | -34.1471 | -2.5544
1.465 | -34.1473 | -2.5653
1.4711 | -34.1478 | -2.5767
1.4772 | -34.1483 | -2.59
1.4833 | -34.1487 | -2.6031
1.4894 | -34.1492 | -2.6153
1.4955 | -34.1498 | -2.6274
1.5016 | -34.1503 | -2.6399
1.5078 | -34.1507 | -2.6538
1.5139 | -34.1511 | -2.6669
1.52 | -34.1515 | -2.6795
1.5261 | -34.1518 | -2.6921
1.5322 | -34.1521 | -2.7062
1.5383 | -34.1525 | -2.7187
1.5444 | -34.153 | -2.7314
1.5505 | -34.1535 | -2.7448
1.5566 | -34.154 | -2.7578
1.5627 | -34.1544 | -2.7716
1.5688 | -34.155 | -2.7852
1.5749 | -34.1554 | -2.7978
1.581 | -34.1557 | -2.8112
1.5871 | -34.156 | -2.8249
1.5932 | -34.1564 | -2.8378
1.5993 | -34.1568 | -2.852
1.6054 | -34.1572 | -2.8653
1.6115 | -34.1576 | -2.8794
1.6176 | -34.1579 | -2.892
1.6237 | -34.1581 | -2.9056
1.6298 | -34.1583 | -2.9182
1.6359 | -34.1587 | -2.931
1.642 | -34.1591 | -2.9429
1.6482 | -34.1593 | -2.9559
1.6543 | -34.1596 | -2.9696
1.6604 | -34.1599 | -2.9817
1.6665 | -34.16 | -2.9942
1.6726 | -34.1602 | -3.0067
1.6787 | -34.1606 | -3.0193
1.6848 | -34.1605 | -3.0319
1.6909 | -34.1607 | -3.0446
1.697 | -34.1609 | -3.0568
1.7031 | -34.161 | -3.0694
1.7092 | -34.1609 | -3.0816
1.7153 | -34.1611 | -3.0936
1.7214 | -34.1609 | -3.1059
1.7275 | -34.1608 | -3.1185
1.7336 | -34.1607 | -3.1315
1.7397 | -34.1609 | -3.1439
1.7458 | -34.1609 | -3.1561
1.7519 | -34.1609 | -3.1683
1.758 | -34.1609 | -3.1803
1.7641 | -34.161 | -3.192
1.7702 | -34.161 | -3.205
1.7763 | -34.1609 | -3.2165
1.7824 | -34.1609 | -3.229
1.7885 | -34.1608 | -3.2415
1.7947 | -34.1609 | -3.2536
1.8008 | -34.1609 | -3.2649
1.8069 | -34.1609 | -3.2765
1.813 | -34.161 | -3.2887
1.8191 | -34.1609 | -3.3015
1.8252 | -34.161 | -3.3144
1.8313 | -34.1611 | -3.3267
1.8374 | -34.1612 | -3.3386
1.8435 | -34.1613 | -3.3506
1.8496 | -34.1613 | -3.3638
1.8557 | -34.1615 | -3.3766
1.8618 | -34.1617 | -3.389
1.8679 | -34.1618 | -3.4011
1.874 | -34.1619 | -3.4132
1.8801 | -34.162 | -3.4257
1.8862 | -34.1623 | -3.4377
1.8923 | -34.1624 | -3.45
1.8984 | -34.1625 | -3.4622
1.9045 | -34.1625 | -3.474
1.9106 | -34.1626 | -3.4863
1.9167 | -34.1626 | -3.4986
1.9228 | -34.1627 | -3.5106
1.9289 | -34.1628 | -3.5226
1.9351 | -34.163 | -3.5337
1.9412 | -34.1632 | -3.5451
1.9473 | -34.1635 | -3.5573
1.9534 | -34.1637 | -3.5688
1.9595 | -34.1638 | -3.5808
1.9656 | -34.1641 | -3.5933
1.9717 | -34.1643 | -3.6052
1.9778 | -34.1646 | -3.6167
1.9839 | -34.1649 | -3.6284
1.99 | -34.165 | -3.6396
1.9961 | -34.1651 | -3.6519
2.0022 | -34.1654 | -3.6635
2.0083 | -34.1654 | -3.6749
2.0144 | -34.1656 | -3.6865
2.0205 | -34.1659 | -3.6984
2.0266 | -34.1663 | -3.71
2.0327 | -34.1665 | -3.7218
2.0388 | -34.1669 | -3.7336
2.0449 | -34.1671 | -3.7451
2.051 | -34.1674 | -3.7567
2.0571 | -34.1677 | -3.7682
2.0632 | -34.1679 | -3.7798
2.0693 | -34.1681 | -3.7922
2.0754 | -34.1684 | -3.8037
2.0816 | -34.1687 | -3.816
2.0877 | -34.1689 | -3.8277
2.0938 | -34.1692 | -3.84
2.0999 | -34.1694 | -3.8518
2.106 | -34.1696 | -3.8636
2.1121 | -34.1698 | -3.8757
2.1182 | -34.1701 | -3.8867
2.1243 | -34.1703 | -3.8976
2.1304 | -34.1705 | -3.909
2.1365 | -34.1708 | -3.9209
2.1426 | -34.1711 | -3.933
2.1487 | -34.1714 | -3.9444
2.1548 | -34.1718 | -3.9564
2.1609 | -34.1722 | -3.9679
2.167 | -34.1726 | -3.9793
2.1731 | -34.1729 | -3.9915
2.1792 | -34.1732 | -4.0027
2.1853 | -34.1736 | -4.0148
2.1914 | -34.1738 | -4.0263
2.1975 | -34.1742 | -4.0381
2.2036 | -34.1745 | -4.0503
2.2097 | -34.1747 | -4.062
2.2158 | -34.1751 | -4.0728
2.222 | -34.1754 | -4.0834
2.2281 | -34.1756 | -4.095
2.2342 | -34.176 | -4.1065
2.2403 | -34.1762 | -4.1173
2.2464 | -34.1765 | -4.1282
2.2525 | -34.1767 | -4.1397
2.2586 | -34.1771 | -4.1505
2.2647 | -34.1774 | -4.1611
2.2708 | -34.1777 | -4.1724
2.2769 | -34.1781 | -4.1833
2.283 | -34.1782 | -4.1946
2.2891 | -34.1785 | -4.2057
2.2952 | -34.1787 | -4.2163
2.3013 | -34.1789 | -4.2274
2.3074 | -34.1792 | -4.238
2.3135 | -34.1794 | -4.2495
2.3196 | -34.1796 | -4.2599
2.3257 | -34.1799 | -4.2707
2.3318 | -34.18 | -4.282
2.3379 | -34.1802 | -4.2929
2.344 | -34.1804 | -4.3045
2.3501 | -34.1806 | -4.3153
2.3562 | -34.1807 | -4.3259
2.3623 | -34.1809 | -4.3364
2.3685 | -34.181 | -4.3471
2.3746 | -34.1812 | -4.3585
2.3807 | -34.1814 | -4.3691
2.3868 | -34.1816 | -4.38
2.3929 | -34.1818 | -4.3908
2.399 | -34.1818 | -4.4015
2.4051 | -34.1821 | -4.4121
2.4112 | -34.1823 | -4.4229
2.4173 | -34.1825 | -4.4337
2.4234 | -34.1826 | -4.4447
2.4295 | -34.1828 | -4.4557
2.4356 | -34.1829 | -4.4667
2.4417 | -34.1831 | -4.478
2.4478 | -34.1832 | -4.489
2.4539 | -34.1834 | -4.4994
2.46 | -34.1835 | -4.5109
2.4661 | -34.1836 | -4.5221
2.4722 | -34.1838 | -4.5325
2.4783 | -34.1839 | -4.5431
2.4844 | -34.184 | -4.5542
2.4905 | -34.1841 | -4.5654
2.4966 | -34.1842 | -4.5763
2.5027 | -34.1843 | -4.5869
2.5089 | -34.1845 | -4.5979
2.515 | -34.1845 | -4.6089
2.5211 | -34.1846 | -4.6198
2.5272 | -34.1846 | -4.6308
2.5333 | -34.1848 | -4.6419
2.5394 | -34.1849 | -4.6528
2.5455 | -34.1851 | -4.6643
2.5516 | -34.1852 | -4.6748
2.5577 | -34.1854 | -4.6864
2.5638 | -34.1855 | -4.6982
2.5699 | -34.1856 | -4.7098
2.576 | -34.1856 | -4.7214
2.5821 | -34.1857 | -4.7328
2.5882 | -34.1858 | -4.7441
2.5943 | -34.186 | -4.7551
2.6004 | -34.1862 | -4.7662
2.6065 | -34.1863 | -4.7767
2.6126 | -34.1864 | -4.7873
2.6187 | -34.1865 | -4.7979
2.6248 | -34.1866 | -4.8088
2.6309 | -34.1868 | -4.8203
2.637 | -34.187 | -4.832
2.6431 | -34.1872 | -4.8426
2.6492 | -34.1873 | -4.8541
2.6554 | -34.1875 | -4.8651
2.6615 | -34.1876 | -4.8762
2.6676 | -34.1878 | -4.8878
2.6737 | -34.1881 | -4.899
2.6798 | -34.1884 | -4.9102
2.6859 | -34.1885 | -4.9216
2.692 | -34.1887 | -4.9326
2.6981 | -34.1889 | -4.9439
2.7042 | -34.1891 | -4.9551
2.7103 | -34.1892 | -4.9659
2.7164 | -34.1895 | -4.977
2.7225 | -34.1898 | -4.9884
2.7286 | -34.19 | -4.9996
2.7347 | -34.1902 | -5.0102
2.7408 | -34.1905 | -5.0205
2.7469 | -34.1907 | -5.0319
2.753 | -34.191 | -5.0426
2.7591 | -34.1913 | -5.054
2.7652 | -34.1915 | -5.0656
2.7713 | -34.1916 | -5.0764
2.7774 | -34.1919 | -5.0874
2.7835 | -34.1921 | -5.0985
2.7896 | -34.1923 | -5.1092
2.7958 | -34.1925 | -5.1201
2.8019 | -34.1928 | -5.1306
2.808 | -34.1931 | -5.1415
2.8141 | -34.1934 | -5.1527
2.8202 | -34.1937 | -5.1637
2.8263 | -34.194 | -5.1747
2.8324 | -34.1943 | -5.1862
2.8385 | -34.1947 | -5.1972
2.8446 | -34.1951 | -5.208
2.8507 | -34.1954 | -5.219
2.8568 | -34.1958 | -5.2295
2.8629 | -34.196 | -5.2403
2.869 | -34.1963 | -5.2515
2.8751 | -34.1966 | -5.2623
2.8812 | -34.1971 | -5.2737
2.8873 | -34.1974 | -5.2848
2.8934 | -34.1977 | -5.2959
2.8995 | -34.198 | -5.3071
2.9056 | -34.1983 | -5.3184
2.9117 | -34.1986 | -5.3295
2.9178 | -34.199 | -5.3403
2.9239 | -34.1994 | -5.352
2.93 | -34.1998 | -5.3635
2.9361 | -34.2001 | -5.3745
2.9423 | -34.2004 | -5.3853
2.9484 | -34.2008 | -5.3965
2.9545 | -34.2012 | -5.4076
2.9606 | -34.2016 | -5.4191
2.9667 | -34.202 | -5.4303
2.9728 | -34.2023 | -5.4417
2.9789 | -34.2027 | -5.4529
2.985 | -34.2031 | -5.4643
2.9911 | -34.2035 | -5.4758
2.9972 | -34.204 | -5.487
3.0033 | -34.2044 | -5.4986
3.0094 | -34.2048 | -5.5098
3.0155 | -34.2052 | -5.5209
3.0216 | -34.2055 | -5.5322
3.0277 | -34.2059 | -5.544
3.0338 | -34.2063 | -5.5555
3.0399 | -34.2067 | -5.567
3.046 | -34.207 | -5.5786
3.0521 | -34.2073 | -5.5899
3.0582 | -34.2076 | -5.6021
3.0643 | -34.208 | -5.6137
3.0704 | -34.2083 | -5.6256
3.0765 | -34.2086 | -5.6368
3.0827 | -34.2089 | -5.6482
3.0888 | -34.2092 | -5.6596
3.0949 | -34.2096 | -5.6711
3.101 | -34.21 | -5.6827
3.1071 | -34.2103 | -5.6946
3.1132 | -34.2107 | -5.7065
3.1193 | -34.211 | -5.7181
3.1254 | -34.2114 | -5.7295
3.1315 | -34.2118 | -5.7416
3.1376 | -34.2121 | -5.7531
3.1437 | -34.2124 | -5.7648
3.1498 | -34.2128 | -5.777
3.1559 | -34.2131 | -5.7885
3.162 | -34.2133 | -5.8008
3.1681 | -34.2136 | -5.8129
3.1742 | -34.2139 | -5.8246
3.1803 | -34.2143 | -5.8364
3.1864 | -34.2146 | -5.8478
3.1925 | -34.2149 | -5.8595
3.1986 | -34.2152 | -5.8717
3.2047 | -34.2155 | -5.8839
3.2108 | -34.2158 | -5.8951
3.2169 | -34.2161 | -5.907
3.223 | -34.2163 | -5.9189
3.2292 | -34.2166 | -5.9308
3.2353 | -34.2168 | -5.9419
3.2414 | -34.217 | -5.9535
3.2475 | -34.2172 | -5.965
3.2536 | -34.2175 | -5.9766
3.2597 | -34.2178 | -5.9877
3.2658 | -34.218 | -5.9988
3.2719 | -34.2182 | -6.01
3.278 | -34.2185 | -6.0218
3.2841 | -34.2187 | -6.033
3.2902 | -34.219 | -6.0446
3.2963 | -34.2193 | -6.0563
3.3024 | -34.2194 | -6.0675
3.3085 | -34.2197 | -6.0793
3.3146 | -34.2198 | -6.0906
3.3207 | -34.2201 | -6.1022
3.3268 | -34.2203 | -6.1138
3.3329 | -34.2205 | -6.1248
3.339 | -34.2207 | -6.1358
3.3451 | -34.2209 | -6.1472
3.3512 | -34.221 | -6.158
3.3573 | -34.2213 | -6.1694
3.3634 | -34.2215 | -6.1807
3.3696 | -34.2217 | -6.1923
3.3757 | -34.2218 | -6.2032
3.3818 | -34.2221 | -6.2143
3.3879 | -34.2222 | -6.2255
3.394 | -34.2224 | -6.237
3.4001 | -34.2226 | -6.248
3.4062 | -34.2227 | -6.2591
3.4123 | -34.2228 | -6.2699
3.4184 | -34.2231 | -6.2813
3.4245 | -34.2233 | -6.2923
3.4306 | -34.2234 | -6.3032
3.4367 | -34.2235 | -6.3139
3.4428 | -34.2237 | -6.3244
3.4489 | -34.2238 | -6.3356
3.455 | -34.224 | -6.3466
3.4611 | -34.2242 | -6.358
3.4672 | -34.2244 | -6.3692
3.4733 | -34.2245 | -6.3804
3.4794 | -34.2247 | -6.3914
3.4855 | -34.2249 | -6.4019
3.4916 | -34.225 | -6.4125
3.4977 | -34.2252 | -6.4239
3.5038 | -34.2253 | -6.435
3.5099 | -34.2255 | -6.4454
3.5161 | -34.2256 | -6.456
3.5222 | -34.2257 | -6.4665
3.5283 | -34.2258 | -6.4775
3.5344 | -34.226 | -6.4884
3.5405 | -34.2262 | -6.4991
3.5466 | -34.2263 | -6.5102
3.5527 | -34.2264 | -6.5211
3.5588 | -34.2266 | -6.5319
3.5649 | -34.2266 | -6.5427
3.571 | -34.2268 | -6.5535
3.5771 | -34.2269 | -6.5648
3.5832 | -34.227 | -6.5748
3.5893 | -34.2273 | -6.5856
3.5954 | -34.2274 | -6.5966
3.6015 | -34.2276 | -6.6076
3.6076 | -34.2278 | -6.6183
3.6137 | -34.228 | -6.6286
3.6198 | -34.2282 | -6.639
3.6259 | -34.2284 | -6.65
3.632 | -34.2286 | -6.6608
3.6381 | -34.2287 | -6.6717
3.6442 | -34.2289 | -6.6824
3.6503 | -34.2291 | -6.6928
3.6565 | -34.2293 | -6.7037
3.6626 | -34.2294 | -6.7144
3.6687 | -34.2297 | -6.7253
3.6748 | -34.2299 | -6.7359
3.6809 | -34.2302 | -6.7467
3.687 | -34.2303 | -6.7578
3.6931 | -34.2306 | -6.7682
3.6992 | -34.2308 | -6.779
3.7053 | -34.231 | -6.7896
3.7114 | -34.2311 | -6.8005
3.7175 | -34.2314 | -6.8113
3.7236 | -34.2316 | -6.8215
3.7297 | -34.2318 | -6.8324
3.7358 | -34.2321 | -6.8435
3.7419 | -34.2323 | -6.8545
3.748 | -34.2325 | -6.8652
3.7541 | -34.2327 | -6.8759
3.7602 | -34.233 | -6.8866
3.7663 | -34.2332 | -6.8972
3.7724 | -34.2334 | -6.908
3.7785 | -34.2337 | -6.9185
3.7846 | -34.2339 | -6.9291
3.7907 | -34.2342 | -6.9402
3.7969 | -34.2345 | -6.9511
3.803 | -34.2347 | -6.9614
3.8091 | -34.2349 | -6.9721
3.8152 | -34.2352 | -6.9828
3.8213 | -34.2355 | -6.9938
3.8274 | -34.2357 | -7.0042
3.8335 | -34.2361 | -7.0151
3.8396 | -34.2364 | -7.0259
3.8457 | -34.2367 | -7.0367
3.8518 | -34.237 | -7.0474
3.8579 | -34.2374 | -7.0581
3.864 | -34.2376 | -7.0691
3.8701 | -34.2379 | -7.0802
3.8762 | -34.2382 | -7.0908
3.8823 | -34.2384 | -7.1017
3.8884 | -34.2388 | -7.1126
3.8945 | -34.2391 | -7.1236
3.9006 | -34.2394 | -7.1348
3.9067 | -34.2397 | -7.146
3.9128 | -34.24 | -7.1569
3.9189 | -34.2403 | -7.1674
3.925 | -34.2406 | -7.1784
3.9311 | -34.2408 | -7.1888
3.9372 | -34.2412 | -7.1998
3.9434 | -34.2414 | -7.2106
3.9495 | -34.2418 | -7.2216
3.9556 | -34.2422 | -7.2322
3.9617 | -34.2425 | -7.2427
3.9678 | -34.2428 | -7.2532
3.9739 | -34.2431 | -7.264
3.98 | -34.2433 | -7.2749
3.9861 | -34.2436 | -7.2856
3.9922 | -34.2439 | -7.2965
3.9983 | -34.2442 | -7.3069
4.0044 | -34.2445 | -7.3173
4.0105 | -34.2448 | -7.3281
4.0166 | -34.2451 | -7.339
4.0227 | -34.2454 | -7.3497
4.0288 | -34.2457 | -7.361
4.0349 | -34.2459 | -7.3715
4.041 | -34.2462 | -7.3823
4.0471 | -34.2464 | -7.3932
4.0532 | -34.2467 | -7.4036
4.0593 | -34.2471 | -7.4146
4.0654 | -34.2474 | -7.4251
4.0715 | -34.2477 | -7.4364
4.0776 | -34.248 | -7.447
4.0838 | -34.2483 | -7.4582
4.0899 | -34.2486 | -7.4687
4.096 | -34.2489 | -7.4794
4.1021 | -34.2491 | -7.49
4.1082 | -34.2494 | -7.5005
4.1143 | -34.2497 | -7.5109
4.1204 | -34.25 | -7.5214
4.1265 | -34.2503 | -7.5318
4.1326 | -34.2506 | -7.5422
4.1387 | -34.2509 | -7.5528
4.1448 | -34.2512 | -7.5636
4.1509 | -34.2514 | -7.5743
4.157 | -34.2516 | -7.5854
4.1631 | -34.2519 | -7.5963
4.1692 | -34.2522 | -7.6068
4.1753 | -34.2525 | -7.6175
4.1814 | -34.2527 | -7.6278
4.1875 | -34.253 | -7.6387
4.1936 | -34.2533 | -7.6496
4.1997 | -34.2535 | -7.6599
4.2058 | -34.2538 | -7.6708
4.2119 | -34.2541 | -7.6813
4.218 | -34.2543 | -7.6917
4.2241 | -34.2545 | -7.7021
4.2303 | -34.2548 | -7.7126
4.2364 | -34.2551 | -7.7232
4.2425 | -34.2553 | -7.7338
4.2486 | -34.2555 | -7.7446
4.2547 | -34.2557 | -7.755
4.2608 | -34.2559 | -7.7657
4.2669 | -34.2561 | -7.776
4.273 | -34.2564 | -7.7867
4.2791 | -34.2566 | -7.7973
4.2852 | -34.2569 | -7.8081
4.2913 | -34.257 | -7.8183
4.2974 | -34.2573 | -7.8292
4.3035 | -34.2575 | -7.84
4.3096 | -34.2577 | -7.8508
4.3157 | -34.258 | -7.8612
4.3218 | -34.2582 | -7.8716
4.3279 | -34.2584 | -7.8822
4.334 | -34.2586 | -7.8927
4.3401 | -34.2589 | -7.9028
4.3462 | -34.2592 | -7.9135
4.3523 | -34.2594 | -7.9236
4.3584 | -34.2596 | -7.9337
4.3645 | -34.2598 | -7.9442
4.3707 | -34.26 | -7.9544
4.3768 | -34.2603 | -7.9648
4.3829 | -34.2605 | -7.9757
4.389 | -34.2607 | -7.9862
4.3951 | -34.2609 | -7.9965
4.4012 | -34.2611 | -8.0073
4.4073 | -34.2613 | -8.0176
4.4134 | -34.2615 | -8.0279
4.4195 | -34.2618 | -8.0383
4.4256 | -34.2621 | -8.0486
4.4317 | -34.2623 | -8.059
4.4378 | -34.2625 | -8.0691
4.4439 | -34.2627 | -8.0792
4.45 | -34.2629 | -8.0897
4.4561 | -34.263 | -8.1001
4.4622 | -34.2633 | -8.1106
4.4683 | -34.2635 | -8.1209
4.4744 | -34.2637 | -8.1316
4.4805 | -34.2639 | -8.1417
4.4866 | -34.2641 | -8.152
4.4927 | -34.2644 | -8.1627
4.4988 | -34.2646 | -8.173
4.5049 | -34.2648 | -8.183
4.511 | -34.2651 | -8.1931
4.5172 | -34.2653 | -8.2034
4.5233 | -34.2655 | -8.2142
4.5294 | -34.2657 | -8.2244
4.5355 | -34.2659 | -8.2348
4.5416 | -34.2662 | -8.2451
4.5477 | -34.2664 | -8.2555
4.5538 | -34.2666 | -8.2654
4.5599 | -34.2668 | -8.2757
4.566 | -34.267 | -8.2861
4.5721 | -34.2673 | -8.2964
4.5782 | -34.2675 | -8.3071
4.5843 | -34.2678 | -8.3175
4.5904 | -34.268 | -8.3279
4.5965 | -34.2683 | -8.3381
4.6026 | -34.2685 | -8.3489
4.6087 | -34.2687 | -8.3592
4.6148 | -34.2689 | -8.3699
4.6209 | -34.2692 | -8.3801
4.627 | -34.2694 | -8.3907
4.6331 | -34.2696 | -8.4009
4.6392 | -34.2698 | -8.4114
4.6453 | -34.27 | -8.4217
4.6514 | -34.2703 | -8.4321
4.6576 | -34.2705 | -8.4429
4.6637 | -34.2707 | -8.4529
4.6698 | -34.2709 | -8.4631
4.6759 | -34.2711 | -8.4739
4.682 | -34.2713 | -8.4841
4.6881 | -34.2715 | -8.4945
4.6942 | -34.2717 | -8.5048
4.7003 | -34.2719 | -8.5152
4.7064 | -34.2722 | -8.5254
4.7125 | -34.2725 | -8.536
4.7186 | -34.2726 | -8.5461
4.7247 | -34.2729 | -8.5566
4.7308 | -34.2732 | -8.5669
4.7369 | -34.2734 | -8.5774
4.743 | -34.2736 | -8.5881
4.7491 | -34.2738 | -8.5988
4.7552 | -34.2741 | -8.6093
4.7613 | -34.2742 | -8.6198
4.7674 | -34.2745 | -8.63
4.7735 | -34.2747 | -8.6409
4.7796 | -34.275 | -8.6515
4.7857 | -34.2752 | -8.6621
4.7918 | -34.2754 | -8.6723
4.7979 | -34.2757 | -8.6826
4.8041 | -34.2759 | -8.693
4.8102 | -34.2762 | -8.7032
4.8163 | -34.2764 | -8.7135
4.8224 | -34.2766 | -8.724
4.8285 | -34.2769 | -8.7341
4.8346 | -34.2772 | -8.7443
4.8407 | -34.2774 | -8.7548
4.8468 | -34.2776 | -8.7655
4.8529 | -34.2779 | -8.7762
4.859 | -34.2781 | -8.7869
4.8651 | -34.2784 | -8.7971
4.8712 | -34.2786 | -8.8073
4.8773 | -34.2788 | -8.8178
4.8834 | -34.2791 | -8.8282
4.8895 | -34.2793 | -8.8384
4.8956 | -34.2795 | -8.8489
4.9017 | -34.2797 | -8.8591
4.9078 | -34.28 | -8.8698
4.9139 | -34.2802 | -8.8798
4.92 | -34.2805 | -8.89
4.9261 | -34.2808 | -8.9
4.9322 | -34.2811 | -8.9105
4.9383 | -34.2813 | -8.9209
4.9445 | -34.2816 | -8.9313
4.9506 | -34.2818 | -8.9415
4.9567 | -34.2821 | -8.9517
4.9628 | -34.2823 | -8.962
4.9689 | -34.2826 | -8.9728
4.975 | -34.2828 | -8.9835
4.9811 | -34.2831 | -8.9936
4.9872 | -34.2834 | -9.0039
4.9933 | -34.2837 | -9.0146
4.9994 | -34.2839 | -9.0254
5.0055 | -34.2842 | -9.0359
5.0116 | -34.2845 | -9.0467
5.0177 | -34.2848 | -9.0574
5.0238 | -34.2851 | -9.0681
5.0299 | -34.2854 | -9.0789
5.036 | -34.2857 | -9.0894
5.0421 | -34.2859 | -9.0998
5.0482 | -34.2862 | -9.1102
5.0543 | -34.2865 | -9.1204
5.0604 | -34.2868 | -9.1309
5.0665 | -34.2871 | -9.141
5.0726 | -34.2873 | -9.1518
5.0787 | -34.2876 | -9.1625
5.0848 | -34.2879 | -9.173
5.091 | -34.2882 | -9.1833
5.0971 | -34.2885 | -9.1934
5.1032 | -34.2888 | -9.2042
5.1093 | -34.2891 | -9.2147
5.1154 | -34.2893 | -9.2249
5.1215 | -34.2895 | -9.2351
5.1276 | -34.2898 | -9.2454
5.1337 | -34.2902 | -9.2555
5.1398 | -34.2904 | -9.2657
5.1459 | -34.2907 | -9.2765
5.152 | -34.291 | -9.2871
5.1581 | -34.2913 | -9.2977
5.1642 | -34.2916 | -9.3083
5.1703 | -34.2919 | -9.3187
5.1764 | -34.2922 | -9.329
5.1825 | -34.2925 | -9.3392
5.1886 | -34.2928 | -9.3494
5.1947 | -34.293 | -9.3599
5.2008 | -34.2933 | -9.3704
5.2069 | -34.2936 | -9.3811
5.213 | -34.2939 | -9.3914
5.2191 | -34.2942 | -9.4019
5.2252 | -34.2945 | -9.4124
5.2314 | -34.2948 | -9.423
5.2375 | -34.2951 | -9.4332
5.2436 | -34.2953 | -9.4438
5.2497 | -34.2956 | -9.454
5.2558 | -34.296 | -9.4648
5.2619 | -34.2962 | -9.4755
5.268 | -34.2966 | -9.4861
5.2741 | -34.2969 | -9.4967
5.2802 | -34.2972 | -9.5075
5.2863 | -34.2975 | -9.5179
5.2924 | -34.2978 | -9.5286
5.2985 | -34.2981 | -9.5392
5.3046 | -34.2984 | -9.5497
5.3107 | -34.2987 | -9.5605
5.3168 | -34.2991 | -9.5711
5.3229 | -34.2993 | -9.5816
5.329 | -34.2996 | -9.5921
5.3351 | -34.3 | -9.6023
5.3412 | -34.3002 | -9.6128
5.3473 | -34.3005 | -9.6234
5.3534 | -34.3009 | -9.6341
5.3595 | -34.3012 | -9.6446
5.3656 | -34.3015 | -9.6552
5.3717 | -34.3018 | -9.666
5.3779 | -34.3021 | -9.6769
5.384 | -34.3024 | -9.6874
5.3901 | -34.3027 | -9.6982
5.3962 | -34.303 | -9.7086
5.4023 | -34.3034 | -9.7193
5.4084 | -34.3037 | -9.7298
5.4145 | -34.304 | -9.7402
5.4206 | -34.3043 | -9.7507
5.4267 | -34.3046 | -9.7616
5.4328 | -34.3049 | -9.7722
5.4389 | -34.3052 | -9.7826
5.445 | -34.3055 | -9.7935
5.4511 | -34.3059 | -9.8038
5.4572 | -34.3062 | -9.8145
5.4633 | -34.3065 | -9.825
5.4694 | -34.3068 | -9.8356
5.4755 | -34.3071 | -9.8459
5.4816 | -34.3074 | -9.8565
5.4877 | -34.3077 | -9.867
5.4938 | -34.308 | -9.8773
5.4999 | -34.3084 | -9.8877
5.506 | -34.3087 | -9.8983
5.5121 | -34.309 | -9.9087
5.5183 | -34.3093 | -9.9191
5.5244 | -34.3096 | -9.9299
5.5305 | -34.31 | -9.9404
5.5366 | -34.3102 | -9.9507
5.5427 | -34.3105 | -9.9614
5.5488 | -34.3108 | -9.972
5.5549 | -34.3111 | -9.9826
5.561 | -34.3115 | -9.9934
5.5671 | -34.3118 | -10.0038
5.5732 | -34.3121 | -10.0146
5.5793 | -34.3124 | -10.0256
5.5854 | -34.3128 | -10.0361
5.5915 | -34.3131 | -10.0469
5.5976 | -34.3134 | -10.0576
5.6037 | -34.3137 | -10.0683
5.6098 | -34.314 | -10.0788
5.6159 | -34.3143 | -10.0892
5.622 | -34.3146 | -10.0999
5.6281 | -34.3149 | -10.1106
5.6342 | -34.3153 | -10.1215
5.6403 | -34.3156 | -10.1321
5.6464 | -34.3159 | -10.1429
5.6525 | -34.3162 | -10.1535
5.6586 | -34.3165 | -10.1639
5.6648 | -34.3169 | -10.1745
5.6709 | -34.3172 | -10.1852
5.677 | -34.3176 | -10.1959
5.6831 | -34.3179 | -10.2065
5.6892 | -34.3182 | -10.2173
5.6953 | -34.3185 | -10.2279
5.7014 | -34.3188 | -10.2387
5.7075 | -34.3191 | -10.2492
5.7136 | -34.3194 | -10.2598
5.7197 | -34.3197 | -10.2704
5.7258 | -34.32 | -10.2808
5.7319 | -34.3203 | -10.2915
5.738 | -34.3207 | -10.302
5.7441 | -34.321 | -10.3126
5.7502 | -34.3214 | -10.3229
5.7563 | -34.3217 | -10.3335
5.7624 | -34.322 | -10.3443
5.7685 | -34.3224 | -10.355
5.7746 | -34.3226 | -10.3659
5.7807 | -34.3229 | -10.3765
5.7868 | -34.3232 | -10.3873
5.7929 | -34.3235 | -10.3978
5.799 | -34.3238 | -10.4083
5.8052 | -34.3241 | -10.4188
5.8113 | -34.3245 | -10.4293
5.8174 | -34.3248 | -10.4398
5.8235 | -34.3251 | -10.4503
5.8296 | -34.3254 | -10.4611
5.8357 | -34.3257 | -10.4717
5.8418 | -34.326 | -10.4824
5.8479 | -34.3263 | -10.4931
5.854 | -34.3266 | -10.5038
5.8601 | -34.3269 | -10.5143
5.8662 | -34.3272 | -10.525
5.8723 | -34.3275 | -10.5355
5.8784 | -34.3278 | -10.546
5.8845 | -34.3281 | -10.5565
5.8906 | -34.3284 | -10.5671
5.8967 | -34.3287 | -10.5775
5.9028 | -34.3291 | -10.588
5.9089 | -34.3295 | -10.5984
5.915 | -34.3297 | -10.6088
5.9211 | -34.33 | -10.6193
5.9272 | -34.3304 | -10.6296
5.9333 | -34.3307 | -10.6401
5.9394 | -34.331 | -10.6507
5.9455 | -34.3313 | -10.661
5.9517 | -34.3317 | -10.6714
5.9578 | -34.3321 | -10.6821
5.9639 | -34.3324 | -10.6923
5.97 | -34.3327 | -10.7029
5.9761 | -34.3331 | -10.7134
5.9822 | -34.3334 | -10.724
5.9883 | -34.3337 | -10.7345
5.9944 | -34.334 | -10.7451
6.0005 | -34.3344 | -10.7558
6.0066 | -34.3347 | -10.7662
6.0127 | -34.335 | -10.7765
6.0188 | -34.3353 | -10.7873
6.0249 | -34.3356 | -10.7979
6.031 | -34.3359 | -10.8082
6.0371 | -34.3363 | -10.8188
6.0432 | -34.3366 | -10.8292
6.0493 | -34.3369 | -10.8395
6.0554 | -34.3373 | -10.8499
6.0615 | -34.3376 | -10.8604
6.0676 | -34.3379 | -10.8711
6.0737 | -34.3383 | -10.8818
6.0798 | -34.3386 | -10.8922
6.0859 | -34.339 | -10.9027
6.0921 | -34.3393 | -10.9129
6.0982 | -34.3396 | -10.9234
6.1043 | -34.3399 | -10.9338
6.1104 | -34.3403 | -10.9445
6.1165 | -34.3406 | -10.9546
6.1226 | -34.341 | -10.965
6.1287 | -34.3413 | -10.9754
6.1348 | -34.3416 | -10.9859
6.1409 | -34.3419 | -10.9963
6.147 | -34.3423 | -11.0067
6.1531 | -34.3426 | -11.0171
6.1592 | -34.3429 | -11.0274
6.1653 | -34.3432 | -11.0377
6.1714 | -34.3435 | -11.0482
6.1775 | -34.3439 | -11.0588
6.1836 | -34.3442 | -11.0692
6.1897 | -34.3446 | -11.0797
6.1958 | -34.3449 | -11.0904
6.2019 | -34.3452 | -11.1009
6.208 | -34.3456 | -11.1114
6.2141 | -34.3459 | -11.1218
6.2202 | -34.3463 | -11.1323
6.2263 | -34.3466 | -11.1426
6.2325 | -34.3469 | -11.1527
6.2386 | -34.3472 | -11.1629
6.2447 | -34.3476 | -11.1736
6.2508 | -34.3479 | -11.184
6.2569 | -34.3483 | -11.1943
6.263 | -34.3486 | -11.2047
6.2691 | -34.3489 | -11.215
6.2752 | -34.3492 | -11.2255
6.2813 | -34.3495 | -11.2359
6.2874 | -34.3499 | -11.2462
6.2935 | -34.3502 | -11.2565
6.2996 | -34.3505 | -11.2669
6.3057 | -34.3508 | -11.2774
6.3118 | -34.3511 | -11.2879
6.3179 | -34.3514 | -11.2982
6.324 | -34.3518 | -11.3086
6.3301 | -34.3521 | -11.3189
6.3362 | -34.3525 | -11.3292
6.3423 | -34.3528 | -11.3397
6.3484 | -34.3531 | -11.3502
6.3545 | -34.3534 | -11.3602
6.3606 | -34.3538 | -11.3703
6.3667 | -34.354 | -11.3805
6.3728 | -34.3544 | -11.391
6.379 | -34.3547 | -11.4013
6.3851 | -34.3551 | -11.4115
6.3912 | -34.3554 | -11.4218
6.3973 | -34.3557 | -11.4321
6.4034 | -34.356 | -11.4423
6.4095 | -34.3563 | -11.4526
6.4156 | -34.3566 | -11.463
6.4217 | -34.3569 | -11.4733
6.4278 | -34.3573 | -11.4837
6.4339 | -34.3576 | -11.4939
6.44 | -34.358 | -11.5042
6.4461 | -34.3583 | -11.5147
6.4522 | -34.3586 | -11.5248
6.4583 | -34.3589 | -11.5353
6.4644 | -34.3592 | -11.5457
6.4705 | -34.3595 | -11.5557
6.4766 | -34.3599 | -11.566
6.4827 | -34.3602 | -11.5763
6.4888 | -34.3605 | -11.5865
6.4949 | -34.3608 | -11.5965
6.501 | -34.3611 | -11.6069
6.5071 | -34.3615 | -11.6169
6.5132 | -34.3618 | -11.6273
6.5194 | -34.3621 | -11.6376
6.5255 | -34.3625 | -11.6476
6.5316 | -34.3627 | -11.6581
6.5377 | -34.3631 | -11.6683
6.5438 | -34.3634 | -11.6784
6.5499 | -34.3637 | -11.6885
6.556 | -34.364 | -11.6988
6.5621 | -34.3643 | -11.7093
6.5682 | -34.3647 | -11.7196
6.5743 | -34.365 | -11.7298
6.5804 | -34.3653 | -11.7401
6.5865 | -34.3657 | -11.7504
6.5926 | -34.366 | -11.7605
6.5987 | -34.3663 | -11.7705
6.6048 | -34.3667 | -11.7808
6.6109 | -34.367 | -11.791
6.617 | -34.3673 | -11.8014
6.6231 | -34.3676 | -11.8116
6.6292 | -34.3679 | -11.8218
6.6353 | -34.3683 | -11.8323
6.6414 | -34.3686 | -11.8424
6.6475 | -34.3689 | -11.8526
6.6536 | -34.3692 | -11.8627
6.6597 | -34.3696 | -11.8727
6.6659 | -34.3699 | -11.8829
6.672 | -34.3702 | -11.893
6.6781 | -34.3705 | -11.9032
6.6842 | -34.3709 | -11.9133
6.6903 | -34.3713 | -11.9235
6.6964 | -34.3716 | -11.9337
6.7025 | -34.3719 | -11.9437
6.7086 | -34.3722 | -11.9539
6.7147 | -34.3725 | -11.9642
6.7208 | -34.3728 | -11.9745
6.7269 | -34.3732 | -11.9846
6.733 | -34.3735 | -11.9946
6.7391 | -34.3738 | -12.0047
6.7452 | -34.3742 | -12.0147
6.7513 | -34.3745 | -12.0246
6.7574 | -34.3748 | -12.0348
6.7635 | -34.3751 | -12.0448
6.7696 | -34.3755 | -12.0548
6.7757 | -34.3758 | -12.065
6.7818 | -34.3762 | -12.0751
6.7879 | -34.3765 | -12.0851
6.794 | -34.3768 | -12.095
6.8001 | -34.3771 | -12.1051
6.8063 | -34.3774 | -12.1152
6.8124 | -34.3778 | -12.1255
6.8185 | -34.3781 | -12.1358
6.8246 | -34.3784 | -12.1457
6.8307 | -34.3787 | -12.1557
6.8368 | -34.3791 | -12.1661
6.8429 | -34.3794 | -12.1762
6.849 | -34.3797 | -12.1864
6.8551 | -34.3801 | -12.1964
6.8612 | -34.3804 | -12.2064
6.8673 | -34.3807 | -12.2165
6.8734 | -34.3811 | -12.2267
6.8795 | -34.3814 | -12.2368
6.8856 | -34.3817 | -12.2469
6.8917 | -34.3821 | -12.2569
6.8978 | -34.3824 | -12.2668
6.9039 | -34.3827 | -12.2766
6.91 | -34.3831 | -12.2867
6.9161 | -34.3834 | -12.2966
6.9222 | -34.3837 | -12.3066
6.9283 | -34.3841 | -12.3164
6.9344 | -34.3844 | -12.3265
6.9405 | -34.3848 | -12.3365
6.9466 | -34.3852 | -12.3468
6.9528 | -34.3855 | -12.357
6.9589 | -34.3859 | -12.3672
6.965 | -34.3862 | -12.3774
6.9711 | -34.3866 | -12.3875
6.9772 | -34.3869 | -12.3976
6.9833 | -34.3873 | -12.4076
6.9894 | -34.3876 | -12.4179
6.9955 | -34.3879 | -12.4277
7.0016 | -34.3883 | -12.4377
7.0077 | -34.3886 | -12.4478
7.0138 | -34.389 | -12.4577
7.0199 | -34.3893 | -12.4677
7.026 | -34.3896 | -12.4777
7.0321 | -34.39 | -12.4875
7.0382 | -34.3904 | -12.4972
7.0443 | -34.3907 | -12.5074
7.0504 | -34.391 | -12.5176
7.0565 | -34.3914 | -12.5276
7.0626 | -34.3917 | -12.5374
7.0687 | -34.3921 | -12.5469
7.0748 | -34.3925 | -12.5567
7.0809 | -34.3928 | -12.5667
7.087 | -34.3932 | -12.5767
7.0932 | -34.3935 | -12.5868
7.0993 | -34.3938 | -12.5967
7.1054 | -34.3942 | -12.6065
7.1115 | -34.3945 | -12.6163
7.1176 | -34.3949 | -12.6262
7.1237 | -34.3952 | -12.6361
7.1298 | -34.3956 | -12.646
7.1359 | -34.3959 | -12.656
7.142 | -34.3962 | -12.6657
7.1481 | -34.3966 | -12.6756
7.1542 | -34.3969 | -12.6856
7.1603 | -34.3973 | -12.6955
7.1664 | -34.3976 | -12.7054
7.1725 | -34.398 | -12.7154
7.1786 | -34.3984 | -12.7253
7.1847 | -34.3987 | -12.7353
7.1908 | -34.3991 | -12.7454
7.1969 | -34.3994 | -12.7556
7.203 | -34.3997 | -12.7653
7.2091 | -34.4001 | -12.7752
7.2152 | -34.4005 | -12.7851
7.2213 | -34.4008 | -12.7949
7.2274 | -34.4012 | -12.8048
7.2335 | -34.4015 | -12.8147
7.2397 | -34.4019 | -12.8245
7.2458 | -34.4023 | -12.8342
7.2519 | -34.4026 | -12.8442
7.258 | -34.403 | -12.8543
7.2641 | -34.4033 | -12.8642
7.2702 | -34.4036 | -12.8743
7.2763 | -34.404 | -12.8842
7.2824 | -34.4044 | -12.8943
7.2885 | -34.4047 | -12.9043
7.2946 | -34.4051 | -12.9142
7.3007 | -34.4054 | -12.9242
7.3068 | -34.4058 | -12.9344
7.3129 | -34.4061 | -12.9445
7.319 | -34.4065 | -12.9544
7.3251 | -34.4068 | -12.9641
7.3312 | -34.4072 | -12.9741
7.3373 | -34.4076 | -12.9838
7.3434 | -34.408 | -12.9936
7.3495 | -34.4083 | -13.0034
7.3556 | -34.4087 | -13.0134
7.3617 | -34.4091 | -13.0235
7.3678 | -34.4094 | -13.0333
7.3739 | -34.4097 | -13.0435
7.3801 | -34.4101 | -13.0537
7.3862 | -34.4105 | -13.0635
7.3923 | -34.4108 | -13.0736
7.3984 | -34.4112 | -13.0833
7.4045 | -34.4116 | -13.0932
7.4106 | -34.412 | -13.103
7.4167 | -34.4123 | -13.113
7.4228 | -34.4127 | -13.1231
7.4289 | -34.4131 | -13.133
7.435 | -34.4134 | -13.1427
7.4411 | -34.4138 | -13.1527
7.4472 | -34.4141 | -13.1627
7.4533 | -34.4145 | -13.1724
7.4594 | -34.4149 | -13.1826
7.4655 | -34.4153 | -13.1923
7.4716 | -34.4156 | -13.2024
7.4777 | -34.416 | -13.2123
7.4838 | -34.4163 | -13.2221
7.4899 | -34.4167 | -13.232
7.496 | -34.4171 | -13.2417
7.5021 | -34.4174 | -13.2516
7.5082 | -34.4178 | -13.2615
7.5143 | -34.4182 | -13.2715
7.5204 | -34.4185 | -13.2815
7.5266 | -34.4189 | -13.2914
7.5327 | -34.4193 | -13.3012
7.5388 | -34.4196 | -13.3114
7.5449 | -34.42 | -13.3215
7.551 | -34.4204 | -13.3313
7.5571 | -34.4207 | -13.3413
7.5632 | -34.4211 | -13.3513
7.5693 | -34.4215 | -13.3611
7.5754 | -34.4219 | -13.3709
7.5815 | -34.4222 | -13.3809
7.5876 | -34.4226 | -13.391
7.5937 | -34.423 | -13.4011
7.5998 | -34.4234 | -13.411
7.6059 | -34.4238 | -13.4211
7.612 | -34.4241 | -13.431
7.6181 | -34.4245 | -13.4408
7.6242 | -34.4249 | -13.451
7.6303 | -34.4253 | -13.4612
7.6364 | -34.4256 | -13.4712
7.6425 | -34.426 | -13.4812
7.6486 | -34.4263 | -13.4911
7.6547 | -34.4267 | -13.501
7.6608 | -34.4271 | -13.5109
7.667 | -34.4274 | -13.521
7.6731 | -34.4278 | -13.531
7.6792 | -34.4282 | -13.5412
7.6853 | -34.4286 | -13.5512
7.6914 | -34.429 | -13.5613
7.6975 | -34.4293 | -13.5714
7.7036 | -34.4297 | -13.5814
7.7097 | -34.4301 | -13.5915
7.7158 | -34.4305 | -13.6016
7.7219 | -34.4309 | -13.6114
7.728 | -34.4312 | -13.6215
7.7341 | -34.4316 | -13.6316
7.7402 | -34.432 | -13.6416
7.7463 | -34.4324 | -13.6517
7.7524 | -34.4327 | -13.6619
7.7585 | -34.4331 | -13.672
7.7646 | -34.4335 | -13.6819
7.7707 | -34.4338 | -13.6922
7.7768 | -34.4342 | -13.7023
7.7829 | -34.4346 | -13.7126
7.789 | -34.4349 | -13.7228
7.7951 | -34.4353 | -13.7328
7.8012 | -34.4357 | -13.7429
7.8073 | -34.4361 | -13.7531
7.8135 | -34.4365 | -13.7631
7.8196 | -34.4369 | -13.7734
7.8257 | -34.4372 | -13.7836
7.8318 | -34.4376 | -13.7937
7.8379 | -34.438 | -13.8037
7.844 | -34.4383 | -13.8137
7.8501 | -34.4387 | -13.8238
7.8562 | -34.4391 | -13.8341
7.8623 | -34.4395 | -13.8441
7.8684 | -34.4399 | -13.8542
7.8745 | -34.4402 | -13.8643
7.8806 | -34.4406 | -13.8745
7.8867 | -34.441 | -13.8843
7.8928 | -34.4414 | -13.8942
7.8989 | -34.4418 | -13.9042
7.905 | -34.4421 | -13.9144
7.9111 | -34.4425 | -13.9245
7.9172 | -34.4429 | -13.9345
7.9233 | -34.4432 | -13.9447
7.9294 | -34.4436 | -13.9548
7.9355 | -34.444 | -13.9647
7.9416 | -34.4443 | -13.9748
7.9477 | -34.4447 | -13.9847
7.9539 | -34.4451 | -13.9947
7.96 | -34.4455 | -14.0049
7.9661 | -34.4459 | -14.015
7.9722 | -34.4463 | -14.025
7.9783 | -34.4467 | -14.0353
7.9844 | -34.4471 | -14.0453
7.9905 | -34.4475 | -14.0556
7.9966 | -34.4478 | -14.0658
8.0027 | -34.4482 | -14.0759
8.0088 | -34.4486 | -14.0859
8.0149 | -34.4489 | -14.096
8.021 | -34.4493 | -14.1062
8.0271 | -34.4497 | -14.1166
8.0332 | -34.4501 | -14.1266
8.0393 | -34.4505 | -14.1369
8.0454 | -34.4508 | -14.1472
8.0515 | -34.4513 | -14.1573
8.0576 | -34.4516 | -14.1677
8.0637 | -34.452 | -14.1777
8.0698 | -34.4524 | -14.1876
8.0759 | -34.4527 | -14.1976
8.082 | -34.4531 | -14.2076
8.0881 | -34.4535 | -14.2178
8.0942 | -34.4539 | -14.228
8.1004 | -34.4543 | -14.2382
8.1065 | -34.4547 | -14.2481
8.1126 | -34.455 | -14.2582
8.1187 | -34.4554 | -14.2685
8.1248 | -34.4557 | -14.2784
8.1309 | -34.4561 | -14.2883
8.137 | -34.4565 | -14.2982
8.1431 | -34.4569 | -14.3085
8.1492 | -34.4573 | -14.3186
8.1553 | -34.4577 | -14.3287
8.1614 | -34.458 | -14.339
8.1675 | -34.4584 | -14.3491
8.1736 | -34.4587 | -14.3594
8.1797 | -34.4591 | -14.3695
8.1858 | -34.4595 | -14.3797
8.1919 | -34.4599 | -14.3898
8.198 | -34.4603 | -14.3998
8.2041 | -34.4607 | -14.4099
8.2102 | -34.4611 | -14.42
8.2163 | -34.4614 | -14.4299
8.2224 | -34.4618 | -14.44
8.2285 | -34.4622 | -14.4498
8.2346 | -34.4626 | -14.4599
8.2408 | -34.463 | -14.4698
8.2469 | -34.4634 | -14.48
8.253 | -34.4637 | -14.4899
8.2591 | -34.4641 | -14.5001
8.2652 | -34.4646 | -14.5102
8.2713 | -34.4649 | -14.5204
8.2774 | -34.4653 | -14.5305
8.2835 | -34.4657 | -14.5409
8.2896 | -34.4661 | -14.5508
8.2957 | -34.4665 | -14.561
8.3018 | -34.4669 | -14.5711
8.3079 | -34.4673 | -14.581
8.314 | -34.4676 | -14.5911
8.3201 | -34.4681 | -14.6011
8.3262 | -34.4684 | -14.6112
8.3323 | -34.4688 | -14.6213
8.3384 | -34.4692 | -14.6315
8.3445 | -34.4696 | -14.6415
8.3506 | -34.47 | -14.6517
8.3567 | -34.4704 | -14.6619
8.3628 | -34.4707 | -14.6719
8.3689 | -34.4711 | -14.6819
8.375 | -34.4715 | -14.6921
8.3812 | -34.4719 | -14.702
8.3873 | -34.4723 | -14.7121
8.3934 | -34.4726 | -14.7223
8.3995 | -34.473 | -14.7322
8.4056 | -34.4734 | -14.7423
8.4117 | -34.4738 | -14.7524
8.4178 | -34.4741 | -14.7626
8.4239 | -34.4745 | -14.7725
8.43 | -34.4749 | -14.7826
8.4361 | -34.4752 | -14.7927
8.4422 | -34.4757 | -14.8027
8.4483 | -34.476 | -14.8127
8.4544 | -34.4764 | -14.8228
8.4605 | -34.4768 | -14.8327
8.4666 | -34.4772 | -14.843
8.4727 | -34.4776 | -14.8531
8.4788 | -34.478 | -14.8632
8.4849 | -34.4783 | -14.8732
8.491 | -34.4787 | -14.8833
8.4971 | -34.4791 | -14.8931
8.5032 | -34.4795 | -14.903
8.5093 | -34.4799 | -14.9129
8.5154 | -34.4803 | -14.9231
8.5215 | -34.4807 | -14.933
8.5277 | -34.481 | -14.9431
8.5338 | -34.4814 | -14.953
8.5399 | -34.4818 | -14.9631
8.546 | -34.4822 | -14.9729
8.5521 | -34.4826 | -14.9829
8.5582 | -34.4829 | -14.9929
8.5643 | -34.4833 | -15.0027
8.5704 | -34.4837 | -15.0124
8.5765 | -34.4841 | -15.0226
8.5826 | -34.4845 | -15.0323
8.5887 | -34.4849 | -15.0422
8.5948 | -34.4853 | -15.0522
8.6009 | -34.4857 | -15.062
8.607 | -34.4861 | -15.0722
8.6131 | -34.4865 | -15.0823
8.6192 | -34.4869 | -15.0921
8.6253 | -34.4873 | -15.102
8.6314 | -34.4876 | -15.112
8.6375 | -34.488 | -15.122
8.6436 | -34.4884 | -15.132
8.6497 | -34.4888 | -15.1418
8.6558 | -34.4891 | -15.1518
8.6619 | -34.4895 | -15.1617
8.6681 | -34.4899 | -15.1719
8.6742 | -34.4903 | -15.1818
8.6803 | -34.4906 | -15.1917
8.6864 | -34.4911 | -15.2014
8.6925 | -34.4914 | -15.2112
8.6986 | -34.4918 | -15.2212
8.7047 | -34.4922 | -15.2314
8.7108 | -34.4925 | -15.2413
8.7169 | -34.4929 | -15.2513
8.723 | -34.4933 | -15.261
8.7291 | -34.4937 | -15.2708
8.7352 | -34.4941 | -15.281
8.7413 | -34.4944 | -15.291
8.7474 | -34.4948 | -15.3007
8.7535 | -34.4952 | -15.3106
8.7596 | -34.4956 | -15.3205
8.7657 | -34.496 | -15.3306
8.7718 | -34.4964 | -15.3406
8.7779 | -34.4968 | -15.3506
8.784 | -34.4972 | -15.3605
8.7901 | -34.4976 | -15.3704
8.7962 | -34.498 | -15.3801
8.8023 | -34.4983 | -15.39
8.8084 | -34.4987 | -15.3999
8.8146 | -34.4991 | -15.4097
8.8207 | -34.4995 | -15.4196
8.8268 | -34.4999 | -15.4292
8.8329 | -34.5003 | -15.439
8.839 | -34.5007 | -15.449
8.8451 | -34.5011 | -15.4589
8.8512 | -34.5015 | -15.4687
8.8573 | -34.5018 | -15.4786
8.8634 | -34.5023 | -15.4883
8.8695 | -34.5026 | -15.4984
8.8756 | -34.503 | -15.5083
8.8817 | -34.5034 | -15.518
8.8878 | -34.5038 | -15.5281
8.8939 | -34.5041 | -15.5379
8.9 | -34.5046 | -15.5478
8.9061 | -34.5049 | -15.5577
8.9122 | -34.5053 | -15.5675
8.9183 | -34.5057 | -15.5775
8.9244 | -34.5061 | -15.5874
8.9305 | -34.5066 | -15.5971
8.9366 | -34.5069 | -15.607
8.9427 | -34.5073 | -15.617
8.9488 | -34.5077 | -15.6267
8.955 | -34.5081 | -15.6365
8.9611 | -34.5085 | -15.6465
8.9672 | -34.5089 | -15.6563
8.9733 | -34.5093 | -15.666
8.9794 | -34.5098 | -15.6757
8.9855 | -34.5102 | -15.6856
8.9916 | -34.5105 | -15.6956
8.9977 | -34.5109 | -15.7054
9.0038 | -34.5114 | -15.7152
9.0099 | -34.5118 | -15.7251
9.016 | -34.5121 | -15.735
9.0221 | -34.5125 | -15.7446
9.0282 | -34.5129 | -15.7543
9.0343 | -34.5133 | -15.7642
9.0404 | -34.5138 | -15.774
9.0465 | -34.5142 | -15.784
9.0526 | -34.5146 | -15.7938
9.0587 | -34.515 | -15.8037
9.0648 | -34.5154 | -15.8134
9.0709 | -34.5158 | -15.8232
9.077 | -34.5163 | -15.8331
9.0831 | -34.5166 | -15.8431
9.0892 | -34.5171 | -15.853
9.0953 | -34.5175 | -15.8629
9.1015 | -34.5179 | -15.8727
9.1076 | -34.5183 | -15.8823
9.1137 | -34.5187 | -15.8921
9.1198 | -34.5191 | -15.9019
9.1259 | -34.5195 | -15.9117
9.132 | -34.5199 | -15.9217
9.1381 | -34.5204 | -15.9313
9.1442 | -34.5208 | -15.9408
9.1503 | -34.5212 | -15.9507
9.1564 | -34.5216 | -15.9605
9.1625 | -34.522 | -15.9703
9.1686 | -34.5224 | -15.98
9.1747 | -34.5229 | -15.9898
9.1808 | -34.5233 | -15.9996
9.1869 | -34.5237 | -16.0093
9.193 | -34.5241 | -16.0191
9.1991 | -34.5245 | -16.029
9.2052 | -34.5249 | -16.0387
9.2113 | -34.5253 | -16.0485
9.2174 | -34.5258 | -16.0583
9.2235 | -34.5262 | -16.0681
9.2296 | -34.5267 | -16.0778
9.2357 | -34.5271 | -16.0878
9.2419 | -34.5275 | -16.0976
9.248 | -34.528 | -16.1073
9.2541 | -34.5284 | -16.1172
9.2602 | -34.5288 | -16.1271
9.2663 | -34.5292 | -16.1366
9.2724 | -34.5296 | -16.1464
9.2785 | -34.53 | -16.1562
9.2846 | -34.5305 | -16.1659
9.2907 | -34.5309 | -16.1757
9.2968 | -34.5313 | -16.1856
9.3029 | -34.5317 | -16.1956
9.309 | -34.5321 | -16.2053
9.3151 | -34.5325 | -16.215
9.3212 | -34.533 | -16.2246
9.3273 | -34.5334 | -16.2342
9.3334 | -34.5338 | -16.2439
9.3395 | -34.5342 | -16.2536
9.3456 | -34.5346 | -16.2632
9.3517 | -34.535 | -16.2731
9.3578 | -34.5354 | -16.2829
9.3639 | -34.5358 | -16.2925
9.37 | -34.5362 | -16.3023
9.3761 | -34.5367 | -16.3122
9.3822 | -34.5371 | -16.322
9.3884 | -34.5375 | -16.3315
9.3945 | -34.538 | -16.3412
9.4006 | -34.5384 | -16.351
9.4067 | -34.5388 | -16.3607
9.4128 | -34.5393 | -16.3706
9.4189 | -34.5397 | -16.3805
9.425 | -34.5401 | -16.39
9.4311 | -34.5406 | -16.3998
9.4372 | -34.541 | -16.4096
9.4433 | -34.5414 | -16.4195
9.4494 | -34.5418 | -16.429
9.4555 | -34.5423 | -16.4389
9.4616 | -34.5427 | -16.4487
9.4677 | -34.5432 | -16.4585
9.4738 | -34.5436 | -16.4682
9.4799 | -34.544 | -16.4779
9.486 | -34.5444 | -16.4877
9.4921 | -34.5448 | -16.4973
9.4982 | -34.5453 | -16.5069
9.5043 | -34.5457 | -16.5166
9.5104 | -34.5461 | -16.5261
9.5165 | -34.5466 | -16.5357
9.5226 | -34.547 | -16.5456
9.5288 | -34.5474 | -16.5552
9.5349 | -34.5478 | -16.565
9.541 | -34.5483 | -16.5744
9.5471 | -34.5487 | -16.5841
9.5532 | -34.5491 | -16.5938
9.5593 | -34.5496 | -16.6036
9.5654 | -34.55 | -16.6133
9.5715 | -34.5504 | -16.623
9.5776 | -34.5508 | -16.6328
9.5837 | -34.5513 | -16.6425
9.5898 | -34.5516 | -16.6523
9.5959 | -34.5521 | -16.662
9.602 | -34.5525 | -16.6715
9.6081 | -34.553 | -16.6811
9.6142 | -34.5534 | -16.6907
9.6203 | -34.5538 | -16.7004
9.6264 | -34.5542 | -16.71
9.6325 | -34.5547 | -16.7199
9.6386 | -34.5551 | -16.7295
9.6447 | -34.5555 | -16.7392
9.6508 | -34.5559 | -16.7488
9.6569 | -34.5563 | -16.7585
9.663 | -34.5568 | -16.7683
9.6691 | -34.5572 | -16.7779
9.6753 | -34.5577 | -16.7876
9.6814 | -34.5581 | -16.7972
9.6875 | -34.5585 | -16.8066
9.6936 | -34.559 | -16.8162
9.6997 | -34.5594 | -16.826
9.7058 | -34.5599 | -16.8357
9.7119 | -34.5603 | -16.8455
9.718 | -34.5607 | -16.8553
9.7241 | -34.5612 | -16.865
9.7302 | -34.5615 | -16.8747
9.7363 | -34.562 | -16.8843
9.7424 | -34.5624 | -16.8939
9.7485 | -34.5629 | -16.9035
9.7546 | -34.5633 | -16.9131
9.7607 | -34.5637 | -16.923
9.7668 | -34.5641 | -16.9327
9.7729 | -34.5646 | -16.9423
9.779 | -34.5651 | -16.9518
9.7851 | -34.5655 | -16.9616
9.7912 | -34.5659 | -16.9713
9.7973 | -34.5664 | -16.9812
9.8034 | -34.5668 | -16.9907
9.8095 | -34.5673 | -17.0006
9.8157 | -34.5677 | -17.0103
9.8218 | -34.5682 | -17.0201
9.8279 | -34.5686 | -17.0297
9.834 | -34.569 | -17.0395
9.8401 | -34.5694 | -17.0491
9.8462 | -34.5699 | -17.0586
9.8523 | -34.5703 | -17.0682
9.8584 | -34.5707 | -17.0779
9.8645 | -34.5711 | -17.0876
9.8706 | -34.5716 | -17.0973
9.8767 | -34.572 | -17.1069
9.8828 | -34.5724 | -17.1165
9.8889 | -34.5728 | -17.126
9.895 | -34.5732 | -17.1357
9.9011 | -34.5737 | -17.1454
9.9072 | -34.5741 | -17.1551
9.9133 | -34.5745 | -17.1647
9.9194 | -34.575 | -17.1743
9.9255 | -34.5754 | -17.184
9.9316 | -34.5758 | -17.1935
9.9377 | -34.5763 | -17.2031
9.9438 | -34.5767 | -17.2128
9.9499 | -34.5771 | -17.2226
9.956 | -34.5775 | -17.2323
9.9622 | -34.578 | -17.2419
9.9683 | -34.5784 | -17.2516
9.9744 | -34.5788 | -17.2614
9.9805 | -34.5793 | -17.271
9.9866 | -34.5797 | -17.2806
9.9927 | -34.5801 | -17.2901
9.9988 | -34.5805 | -17.2998
10.0049 | -34.581 | -17.3095
10.011 | -34.5814 | -17.3193
10.0171 | -34.5819 | -17.3289
10.0232 | -34.5823 | -17.3384
10.0293 | -34.5827 | -17.3481
10.0354 | -34.5832 | -17.3577
10.0415 | -34.5836 | -17.3674
10.0476 | -34.5841 | -17.3772
10.0537 | -34.5845 | -17.387
10.0598 | -34.5849 | -17.3966
10.0659 | -34.5853 | -17.4063
10.072 | -34.5858 | -17.4157
10.0781 | -34.5862 | -17.4253
10.0842 | -34.5867 | -17.4348
10.0903 | -34.5871 | -17.4444
10.0964 | -34.5875 | -17.4538
10.1026 | -34.5879 | -17.4635
10.1087 | -34.5884 | -17.4731
10.1148 | -34.5888 | -17.4828
10.1209 | -34.5892 | -17.4923
10.127 | -34.5896 | -17.5017
10.1331 | -34.5901 | -17.5111
10.1392 | -34.5905 | -17.5206
10.1453 | -34.591 | -17.5302
10.1514 | -34.5914 | -17.5399
10.1575 | -34.5918 | -17.5495
10.1636 | -34.5922 | -17.5588
10.1697 | -34.5926 | -17.5685
10.1758 | -34.5931 | -17.5779
10.1819 | -34.5935 | -17.5875
10.188 | -34.5939 | -17.5971
10.1941 | -34.5943 | -17.6067
10.2002 | -34.5948 | -17.6163
10.2063 | -34.5952 | -17.6259
10.2124 | -34.5956 | -17.6355
10.2185 | -34.5961 | -17.6452
10.2246 | -34.5965 | -17.6547
10.2307 | -34.5969 | -17.6644
10.2368 | -34.5974 | -17.6736
10.2429 | -34.5978 | -17.6833
10.2491 | -34.5982 | -17.6929
10.2552 | -34.5986 | -17.7024
10.2613 | -34.5991 | -17.712
10.2674 | -34.5995 | -17.7217
10.2735 | -34.6 | -17.7314
10.2796 | -34.6004 | -17.7409
10.2857 | -34.6008 | -17.7507
10.2918 | -34.6013 | -17.7605
10.2979 | -34.6017 | -17.77
10.304 | -34.6022 | -17.7795
10.3101 | -34.6026 | -17.7891
10.3162 | -34.6031 | -17.7988
10.3223 | -34.6035 | -17.8083
10.3284 | -34.6039 | -17.8179
10.3345 | -34.6044 | -17.8274
10.3406 | -34.6048 | -17.8369
10.3467 | -34.6052 | -17.8465
10.3528 | -34.6056 | -17.8558
10.3589 | -34.606 | -17.8655
10.365 | -34.6065 | -17.8751
10.3711 | -34.6069 | -17.8848
10.3772 | -34.6073 | -17.8942
10.3833 | -34.6078 | -17.9038
10.3895 | -34.6083 | -17.9133
10.3956 | -34.6087 | -17.9229
10.4017 | -34.6091 | -17.9323
10.4078 | -34.6095 | -17.9419
10.4139 | -34.61 | -17.9516
10.42 | -34.6104 | -17.9612
10.4261 | -34.6109 | -17.9708
10.4322 | -34.6113 | -17.9803
10.4383 | -34.6117 | -17.9898
10.4444 | -34.6121 | -17.9992
10.4505 | -34.6126 | -18.0089
10.4566 | -34.613 | -18.0184
10.4627 | -34.6134 | -18.0279
10.4688 | -34.6139 | -18.0373
10.4749 | -34.6143 | -18.0468
10.481 | -34.6147 | -18.0564
10.4871 | -34.6152 | -18.0658
10.4932 | -34.6156 | -18.0752
10.4993 | -34.616 | -18.0848
10.5054 | -34.6165 | -18.0943
10.5115 | -34.6169 | -18.1039
10.5176 | -34.6173 | -18.1133
10.5237 | -34.6178 | -18.1231
10.5298 | -34.6182 | -18.1324
10.536 | -34.6186 | -18.142
10.5421 | -34.6191 | -18.1515
10.5482 | -34.6195 | -18.1611
10.5543 | -34.62 | -18.1705
10.5604 | -34.6204 | -18.1798
10.5665 | -34.6208 | -18.1893
10.5726 | -34.6212 | -18.1988
10.5787 | -34.6217 | -18.2083
10.5848 | -34.6221 | -18.2177
10.5909 | -34.6226 | -18.2273
10.597 | -34.623 | -18.2369
10.6031 | -34.6234 | -18.2466
10.6092 | -34.6239 | -18.2561
10.6153 | -34.6243 | -18.2658
10.6214 | -34.6248 | -18.2752
10.6275 | -34.6252 | -18.2847
10.6336 | -34.6257 | -18.2942
10.6397 | -34.6261 | -18.3038
10.6458 | -34.6265 | -18.3133
10.6519 | -34.6269 | -18.3229
10.658 | -34.6274 | -18.3323
10.6641 | -34.6278 | -18.3419
10.6702 | -34.6283 | -18.3514
10.6764 | -34.6287 | -18.3611
10.6825 | -34.6292 | -18.3705
10.6886 | -34.6296 | -18.3798
10.6947 | -34.63 | -18.3894
10.7008 | -34.6305 | -18.3989
10.7069 | -34.6309 | -18.4085
10.713 | -34.6314 | -18.4179
10.7191 | -34.6318 | -18.4275
10.7252 | -34.6322 | -18.4367
10.7313 | -34.6326 | -18.4463
10.7374 | -34.6331 | -18.4558
10.7435 | -34.6335 | -18.4654
10.7496 | -34.6339 | -18.4747
10.7557 | -34.6344 | -18.4843
10.7618 | -34.6348 | -18.4935
10.7679 | -34.6353 | -18.503
10.774 | -34.6357 | -18.5125
10.7801 | -34.6362 | -18.5221
10.7862 | -34.6366 | -18.5313
10.7923 | -34.637 | -18.5408
10.7984 | -34.6374 | -18.5502
10.8045 | -34.6379 | -18.5594
10.8106 | -34.6383 | -18.569
10.8168 | -34.6387 | -18.5784
10.8229 | -34.6392 | -18.5879
10.829 | -34.6396 | -18.5973
10.8351 | -34.6401 | -18.6064
10.8412 | -34.6406 | -18.6159
10.8473 | -34.641 | -18.6251
10.8534 | -34.6415 | -18.6346
10.8595 | -34.6419 | -18.644
10.8656 | -34.6424 | -18.6532
10.8717 | -34.6428 | -18.6626
10.8778 | -34.6433 | -18.672
10.8839 | -34.6437 | -18.6814
10.89 | -34.6442 | -18.6908
10.8961 | -34.6446 | -18.7002
10.9022 | -34.6451 | -18.7096
10.9083 | -34.6455 | -18.7191
10.9144 | -34.6459 | -18.7284
10.9205 | -34.6464 | -18.7378
10.9266 | -34.6468 | -18.7471
10.9327 | -34.6473 | -18.7564
10.9388 | -34.6478 | -18.7657
10.9449 | -34.6482 | -18.775
10.951 | -34.6487 | -18.7843
10.9571 | -34.6491 | -18.7937
10.9633 | -34.6496 | -18.8033
10.9694 | -34.65 | -18.8129
10.9755 | -34.6505 | -18.8222
10.9816 | -34.6509 | -18.8317
10.9877 | -34.6514 | -18.8409
10.9938 | -34.6519 | -18.8505
10.9999 | -34.6523 | -18.8598
11.006 | -34.6528 | -18.869
11.0121 | -34.6532 | -18.8784
11.0182 | -34.6537 | -18.8878
11.0243 | -34.6541 | -18.8971
11.0304 | -34.6546 | -18.9064
11.0365 | -34.655 | -18.9156
11.0426 | -34.6555 | -18.9249
11.0487 | -34.656 | -18.9343
11.0548 | -34.6565 | -18.9435
11.0609 | -34.6569 | -18.953
11.067 | -34.6574 | -18.9623
11.0731 | -34.6579 | -18.9716
11.0792 | -34.6583 | -18.9809
11.0853 | -34.6588 | -18.9902
11.0914 | -34.6592 | -18.9996
11.0975 | -34.6597 | -19.009
11.1037 | -34.6601 | -19.0182
11.1098 | -34.6606 | -19.0276
11.1159 | -34.6611 | -19.0369
11.122 | -34.6615 | -19.046
11.1281 | -34.662 | -19.0553
11.1342 | -34.6624 | -19.0647
11.1403 | -34.6629 | -19.0742
11.1464 | -34.6633 | -19.0837
11.1525 | -34.6638 | -19.0929
11.1586 | -34.6642 | -19.1023
11.1647 | -34.6647 | -19.1116
11.1708 | -34.6651 | -19.1208
11.1769 | -34.6655 | -19.1302
11.183 | -34.666 | -19.1392
11.1891 | -34.6665 | -19.1486
11.1952 | -34.667 | -19.1579
11.2013 | -34.6674 | -19.1673
11.2074 | -34.6679 | -19.1765
11.2135 | -34.6684 | -19.186
11.2196 | -34.6688 | -19.1951
11.2257 | -34.6693 | -19.2046
11.2318 | -34.6698 | -19.2137
11.2379 | -34.6702 | -19.2232
11.244 | -34.6707 | -19.2324
11.2502 | -34.6712 | -19.2418
11.2563 | -34.6717 | -19.2511
11.2624 | -34.6722 | -19.2605
11.2685 | -34.6726 | -19.2697
11.2746 | -34.6731 | -19.2788
11.2807 | -34.6736 | -19.2882
11.2868 | -34.674 | -19.2975
11.2929 | -34.6745 | -19.3065
11.299 | -34.675 | -19.3157
11.3051 | -34.6754 | -19.3252
11.3112 | -34.6759 | -19.3345
11.3173 | -34.6764 | -19.3437
11.3234 | -34.6769 | -19.3528
11.3295 | -34.6773 | -19.3623
11.3356 | -34.6778 | -19.3716
11.3417 | -34.6783 | -19.3808
11.3478 | -34.6787 | -19.3902
11.3539 | -34.6791 | -19.3995
11.36 | -34.6796 | -19.4086
11.3661 | -34.6801 | -19.4178
11.3722 | -34.6806 | -19.427
11.3783 | -34.681 | -19.4365
11.3844 | -34.6815 | -19.4459
11.3906 | -34.682 | -19.4552
11.3967 | -34.6825 | -19.4644
11.4028 | -34.683 | -19.4738
11.4089 | -34.6834 | -19.4829
11.415 | -34.6839 | -19.492
11.4211 | -34.6843 | -19.5013
11.4272 | -34.6848 | -19.5105
11.4333 | -34.6853 | -19.5194
11.4394 | -34.6857 | -19.5287
11.4455 | -34.6862 | -19.5379
11.4516 | -34.6867 | -19.5471
11.4577 | -34.6872 | -19.5564
11.4638 | -34.6877 | -19.5655
11.4699 | -34.6882 | -19.5748
11.476 | -34.6887 | -19.5842
11.4821 | -34.6892 | -19.5935
11.4882 | -34.6897 | -19.6028
11.4943 | -34.6901 | -19.612
11.5004 | -34.6906 | -19.6209
11.5065 | -34.6911 | -19.6299
11.5126 | -34.6916 | -19.6391
11.5187 | -34.6921 | -19.6483
11.5248 | -34.6926 | -19.6577
11.5309 | -34.6931 | -19.6669
11.5371 | -34.6935 | -19.6758
11.5432 | -34.694 | -19.6852
11.5493 | -34.6945 | -19.6945
11.5554 | -34.6949 | -19.7037
11.5615 | -34.6954 | -19.713
11.5676 | -34.6959 | -19.7223
11.5737 | -34.6964 | -19.7313
11.5798 | -34.6969 | -19.7403
11.5859 | -34.6974 | -19.7495
11.592 | -34.6979 | -19.7587
11.5981 | -34.6983 | -19.7679
11.6042 | -34.6989 | -19.777
11.6103 | -34.6993 | -19.7861
11.6164 | -34.6998 | -19.7952
11.6225 | -34.7003 | -19.8043
11.6286 | -34.7008 | -19.8134
11.6347 | -34.7012 | -19.8227
11.6408 | -34.7017 | -19.8316
11.6469 | -34.7022 | -19.8408
11.653 | -34.7027 | -19.8498
11.6591 | -34.7032 | -19.859
11.6652 | -34.7037 | -19.8682
11.6713 | -34.7042 | -19.8772
11.6775 | -34.7047 | -19.8863
11.6836 | -34.7051 | -19.8954
11.6897 | -34.7056 | -19.9046
11.6958 | -34.7061 | -19.9138
11.7019 | -34.7066 | -19.9231
11.708 | -34.7071 | -19.9322
11.7141 | -34.7076 | -19.9412
11.7202 | -34.7081 | -19.9504
11.7263 | -34.7085 | -19.9595
11.7324 | -34.709 | -19.9688
11.7385 | -34.7095 | -19.9779
11.7446 | -34.7099 | -19.987
11.7507 | -34.7105 | -19.9958
11.7568 | -34.711 | -20.005
11.7629 | -34.7114 | -20.0142
11.769 | -34.7119 | -20.0232
11.7751 | -34.7124 | -20.0324
11.7812 | -34.7129 | -20.0413
11.7873 | -34.7134 | -20.0504
11.7934 | -34.7139 | -20.0595
11.7995 | -34.7144 | -20.0686
11.8056 | -34.7149 | -20.0777
11.8117 | -34.7153 | -20.0868
11.8178 | -34.7159 | -20.0957
11.824 | -34.7163 | -20.1047
11.8301 | -34.7168 | -20.1138
11.8362 | -34.7173 | -20.1228
11.8423 | -34.7178 | -20.1318
11.8484 | -34.7183 | -20.141
11.8545 | -34.7188 | -20.1499
11.8606 | -34.7193 | -20.1589
11.8667 | -34.7198 | -20.168
11.8728 | -34.7203 | -20.1769
11.8789 | -34.7207 | -20.1859
11.885 | -34.7212 | -20.195
11.8911 | -34.7217 | -20.2039
11.8972 | -34.7223 | -20.2128
11.9033 | -34.7227 | -20.2218
11.9094 | -34.7232 | -20.2308
11.9155 | -34.7237 | -20.2398
11.9216 | -34.7242 | -20.2487
11.9277 | -34.7247 | -20.2577
11.9338 | -34.7252 | -20.2668
11.9399 | -34.7257 | -20.2759
11.946 | -34.7262 | -20.2849
11.9521 | -34.7267 | -20.2938
11.9582 | -34.7272 | -20.3027
11.9644 | -34.7277 | -20.3117
11.9705 | -34.7282 | -20.3207
11.9766 | -34.7287 | -20.3297
11.9827 | -34.7292 | -20.3387
11.9888 | -34.7297 | -20.3476
11.9949 | -34.7302 | -20.3563
12.001 | -34.7307 | -20.3654
12.0071 | -34.7312 | -20.3744
12.0132 | -34.7317 | -20.3832
12.0193 | -34.7322 | -20.3923
12.0254 | -34.7327 | -20.4013
12.0315 | -34.7332 | -20.4099
12.0376 | -34.7337 | -20.4187
12.0437 | -34.7341 | -20.4277
12.0498 | -34.7347 | -20.4366
12.0559 | -34.7351 | -20.4454
12.062 | -34.7357 | -20.4541
12.0681 | -34.7362 | -20.4632
12.0742 | -34.7367 | -20.472
12.0803 | -34.7372 | -20.4811
12.0864 | -34.7377 | -20.4901
12.0925 | -34.7382 | -20.4991
12.0986 | -34.7387 | -20.5079
12.1047 | -34.7392 | -20.5169
12.1109 | -34.7397 | -20.5257
12.117 | -34.7402 | -20.5346
12.1231 | -34.7407 | -20.5435
12.1292 | -34.7412 | -20.5525
12.1353 | -34.7417 | -20.5614
12.1414 | -34.7422 | -20.5704
12.1475 | -34.7427 | -20.5794
12.1536 | -34.7432 | -20.5884
12.1597 | -34.7437 | -20.5974
12.1658 | -34.7442 | -20.6062
12.1719 | -34.7447 | -20.6149
12.178 | -34.7452 | -20.6238
12.1841 | -34.7457 | -20.6328
12.1902 | -34.7462 | -20.6417
12.1963 | -34.7467 | -20.6507
12.2024 | -34.7472 | -20.6593
12.2085 | -34.7477 | -20.6683
12.2146 | -34.7482 | -20.6772
12.2207 | -34.7487 | -20.6861
12.2268 | -34.7492 | -20.695
12.2329 | -34.7497 | -20.7038
12.239 | -34.7503 | -20.7124
12.2451 | -34.7508 | -20.7213
12.2513 | -34.7513 | -20.7302
12.2574 | -34.7518 | -20.7391
12.2635 | -34.7523 | -20.7479
12.2696 | -34.7528 | -20.7567
12.2757 | -34.7533 | -20.7653
12.2818 | -34.7538 | -20.7742
12.2879 | -34.7543 | -20.7831
12.294 | -34.7548 | -20.7918
12.3001 | -34.7553 | -20.8007
12.3062 | -34.7559 | -20.8095
12.3123 | -34.7564 | -20.8181
12.3184 | -34.7569 | -20.8269
12.3245 | -34.7574 | -20.8356
12.3306 | -34.7579 | -20.8444
12.3367 | -34.7584 | -20.8533
12.3428 | -34.7589 | -20.8621
12.3489 | -34.7594 | -20.8706
12.355 | -34.7599 | -20.8796
12.3611 | -34.7604 | -20.8885
12.3672 | -34.7609 | -20.8973
12.3733 | -34.7614 | -20.9062
12.3794 | -34.7619 | -20.9146
12.3855 | -34.7624 | -20.9234
12.3916 | -34.7629 | -20.9322
12.3978 | -34.7634 | -20.9411
12.4039 | -34.764 | -20.9501
12.41 | -34.7645 | -20.9589
12.4161 | -34.765 | -20.9673
12.4222 | -34.7655 | -20.976
12.4283 | -34.766 | -20.9848
12.4344 | -34.7665 | -20.9936
12.4405 | -34.767 | -21.0025
12.4466 | -34.7675 | -21.0113
12.4527 | -34.768 | -21.0199
12.4588 | -34.7685 | -21.0287
12.4649 | -34.769 | -21.0374
12.471 | -34.7695 | -21.0464
12.4771 | -34.77 | -21.0551
12.4832 | -34.7705 | -21.0638
12.4893 | -34.771 | -21.0725
12.4954 | -34.7715 | -21.0811
12.5015 | -34.772 | -21.0899
12.5076 | -34.7725 | -21.0987
12.5137 | -34.773 | -21.1075
12.5198 | -34.7736 | -21.1161
12.5259 | -34.774 | -21.1248
12.532 | -34.7745 | -21.1336
12.5382 | -34.7751 | -21.1424
12.5443 | -34.7756 | -21.1511
12.5504 | -34.7761 | -21.1599
12.5565 | -34.7766 | -21.1685
12.5626 | -34.7771 | -21.1773
12.5687 | -34.7776 | -21.1861
12.5748 | -34.7781 | -21.1949
12.5809 | -34.7786 | -21.2035
12.587 | -34.7791 | -21.2123
12.5931 | -34.7797 | -21.221
12.5992 | -34.7802 | -21.2296
12.6053 | -34.7807 | -21.2383
12.6114 | -34.7812 | -21.2469
12.6175 | -34.7817 | -21.2557
12.6236 | -34.7822 | -21.2645
12.6297 | -34.7827 | -21.2731
12.6358 | -34.7832 | -21.2816
12.6419 | -34.7837 | -21.2904
12.648 | -34.7842 | -21.2992
12.6541 | -34.7847 | -21.3079
12.6602 | -34.7852 | -21.3162
12.6663 | -34.7858 | -21.325
12.6724 | -34.7863 | -21.3336
12.6785 | -34.7868 | -21.3423
12.6847 | -34.7873 | -21.351
12.6908 | -34.7878 | -21.3597
12.6969 | -34.7883 | -21.3682
12.703 | -34.7888 | -21.377
12.7091 | -34.7893 | -21.3857
12.7152 | -34.7898 | -21.3942
12.7213 | -34.7903 | -21.403
12.7274 | -34.7908 | -21.4117
12.7335 | -34.7913 | -21.4201
12.7396 | -34.7918 | -21.4287
12.7457 | -34.7923 | -21.4374
12.7518 | -34.7928 | -21.4461
12.7579 | -34.7933 | -21.4548
12.764 | -34.7938 | -21.4636
12.7701 | -34.7943 | -21.4724
12.7762 | -34.7948 | -21.481
12.7823 | -34.7954 | -21.4897
12.7884 | -34.7959 | -21.4984
12.7945 | -34.7964 | -21.5073
12.8006 | -34.7969 | -21.5159
12.8067 | -34.7974 | -21.5244
12.8128 | -34.7979 | -21.5329
12.8189 | -34.7984 | -21.5415
12.8251 | -34.7989 | -21.5502
12.8312 | -34.7994 | -21.559
12.8373 | -34.7999 | -21.5676
12.8434 | -34.8004 | -21.5763
12.8495 | -34.8009 | -21.5849
12.8556 | -34.8014 | -21.5935
12.8617 | -34.8019 | -21.6021
12.8678 | -34.8024 | -21.611
12.8739 | -34.8029 | -21.6194
12.88 | -34.8034 | -21.6281
12.8861 | -34.8039 | -21.6367
12.8922 | -34.8044 | -21.6454
12.8983 | -34.805 | -21.654
12.9044 | -34.8055 | -21.6625
12.9105 | -34.8059 | -21.6711
12.9166 | -34.8065 | -21.6797
12.9227 | -34.8069 | -21.6884
12.9288 | -34.8074 | -21.697
12.9349 | -34.8079 | -21.7057
12.941 | -34.8085 | -21.7142
12.9471 | -34.809 | -21.7228
12.9532 | -34.8095 | -21.7314
12.9593 | -34.81 | -21.7402
12.9654 | -34.8105 | -21.7489
12.9716 | -34.811 | -21.7577
12.9777 | -34.8115 | -21.7663
12.9838 | -34.812 | -21.775
12.9899 | -34.8125 | -21.7837
12.996 | -34.813 | -21.7923
13.0021 | -34.8136 | -21.801
13.0082 | -34.8141 | -21.8095
13.0143 | -34.8146 | -21.8178
13.0204 | -34.8151 | -21.8264
13.0265 | -34.8156 | -21.8349
13.0326 | -34.8161 | -21.8436
13.0387 | -34.8166 | -21.8522
13.0448 | -34.8171 | -21.8609
13.0509 | -34.8176 | -21.8694
13.057 | -34.8181 | -21.878
13.0631 | -34.8186 | -21.8865
13.0692 | -34.8192 | -21.895
13.0753 | -34.8197 | -21.9038
13.0814 | -34.8202 | -21.9125
13.0875 | -34.8207 | -21.9211
13.0936 | -34.8211 | -21.9297
13.0997 | -34.8217 | -21.9382
13.1058 | -34.8222 | -21.9467
13.112 | -34.8227 | -21.9552
13.1181 | -34.8232 | -21.9637
13.1242 | -34.8237 | -21.9723
13.1303 | -34.8242 | -21.981
13.1364 | -34.8247 | -21.9895
13.1425 | -34.8251 | -21.9981
13.1486 | -34.8256 | -22.0067
13.1547 | -34.8261 | -22.0153
13.1608 | -34.8267 | -22.0239
13.1669 | -34.8272 | -22.0323
13.173 | -34.8277 | -22.041
13.1791 | -34.8282 | -22.0495
13.1852 | -34.8287 | -22.058
13.1913 | -34.8292 | -22.0664
13.1974 | -34.8297 | -22.075
13.2035 | -34.8302 | -22.0835
13.2096 | -34.8307 | -22.0919
13.2157 | -34.8312 | -22.1005
13.2218 | -34.8317 | -22.1089
13.2279 | -34.8322 | -22.1173
13.234 | -34.8327 | -22.1259
13.2401 | -34.8332 | -22.1344
13.2462 | -34.8337 | -22.1429
13.2524 | -34.8342 | -22.1515
13.2585 | -34.8347 | -22.1599
13.2646 | -34.8352 | -22.1685
13.2707 | -34.8357 | -22.1771
13.2768 | -34.8363 | -22.1857
13.2829 | -34.8368 | -22.1943
13.289 | -34.8373 | -22.2027
13.2951 | -34.8378 | -22.211
13.3012 | -34.8383 | -22.2195
13.3073 | -34.8388 | -22.2281
13.3134 | -34.8393 | -22.2367
13.3195 | -34.8398 | -22.2451
13.3256 | -34.8403 | -22.2536
13.3317 | -34.8408 | -22.262
13.3378 | -34.8413 | -22.2707
13.3439 | -34.8418 | -22.2792
13.35 | -34.8423 | -22.288
13.3561 | -34.8428 | -22.2964
13.3622 | -34.8433 | -22.3049
13.3683 | -34.8438 | -22.3133
13.3744 | -34.8443 | -22.3218
13.3805 | -34.8448 | -22.3303
13.3866 | -34.8453 | -22.3387
13.3927 | -34.8458 | -22.3474
13.3989 | -34.8463 | -22.3559
13.405 | -34.8468 | -22.3644
13.4111 | -34.8473 | -22.3729
13.4172 | -34.8478 | -22.3814
13.4233 | -34.8483 | -22.3899
13.4294 | -34.8488 | -22.3986
13.4355 | -34.8492 | -22.4069
13.4416 | -34.8497 | -22.4154
13.4477 | -34.8503 | -22.4238
13.4538 | -34.8508 | -22.4323
13.4599 | -34.8512 | -22.4408
13.466 | -34.8517 | -22.4495
13.4721 | -34.8522 | -22.4578
13.4782 | -34.8527 | -22.4662
13.4843 | -34.8532 | -22.4749
13.4904 | -34.8537 | -22.4833
13.4965 | -34.8541 | -22.4918
13.5026 | -34.8547 | -22.5005
13.5087 | -34.8552 | -22.5089
13.5148 | -34.8557 | -22.5175
13.5209 | -34.8561 | -22.5262
13.527 | -34.8566 | -22.5349
13.5331 | -34.8571 | -22.5432
13.5393 | -34.8576 | -22.5515
13.5454 | -34.8582 | -22.56
13.5515 | -34.8586 | -22.5687
13.5576 | -34.8591 | -22.5773
13.5637 | -34.8596 | -22.5859
13.5698 | -34.8601 | -22.5943
13.5759 | -34.8606 | -22.6025
13.582 | -34.8611 | -22.611
13.5881 | -34.8616 | -22.6194
13.5942 | -34.8621 | -22.628
13.6003 | -34.8626 | -22.6365
13.6064 | -34.8631 | -22.6451
13.6125 | -34.8636 | -22.6536
13.6186 | -34.8641 | -22.6621
13.6247 | -34.8646 | -22.6706
13.6308 | -34.8651 | -22.6792
13.6369 | -34.8656 | -22.6877
13.643 | -34.8661 | -22.6964
13.6491 | -34.8666 | -22.7048
13.6552 | -34.8671 | -22.7132
13.6613 | -34.8676 | -22.7218
13.6674 | -34.868 | -22.7303
13.6735 | -34.8685 | -22.7388
13.6796 | -34.869 | -22.7473
13.6858 | -34.8695 | -22.7557
13.6919 | -34.87 | -22.7642
13.698 | -34.8705 | -22.7727
13.7041 | -34.871 | -22.7812
13.7102 | -34.8715 | -22.7898
13.7163 | -34.872 | -22.798
13.7224 | -34.8725 | -22.8066
13.7285 | -34.873 | -22.815
13.7346 | -34.8735 | -22.8235
13.7407 | -34.874 | -22.832
13.7468 | -34.8745 | -22.8405
13.7529 | -34.875 | -22.8488
13.759 | -34.8755 | -22.8573
13.7651 | -34.876 | -22.8659
13.7712 | -34.8765 | -22.8744
13.7773 | -34.877 | -22.8829
13.7834 | -34.8775 | -22.8914
13.7895 | -34.878 | -22.8997
13.7956 | -34.8785 | -22.9083
13.8017 | -34.879 | -22.9167
13.8078 | -34.8795 | -22.9252
13.8139 | -34.8799 | -22.9335
13.82 | -34.8804 | -22.9418
13.8262 | -34.8809 | -22.9503
13.8323 | -34.8814 | -22.9589
13.8384 | -34.8819 | -22.9674
13.8445 | -34.8824 | -22.9758
13.8506 | -34.8829 | -22.9842
13.8567 | -34.8834 | -22.9927
13.8628 | -34.8839 | -23.0013
13.8689 | -34.8844 | -23.0098
13.875 | -34.8849 | -23.0182
13.8811 | -34.8854 | -23.0267
13.8872 | -34.8859 | -23.0352
13.8933 | -34.8864 | -23.0434
13.8994 | -34.8868 | -23.052
13.9055 | -34.8873 | -23.0606
13.9116 | -34.8878 | -23.0689
13.9177 | -34.8883 | -23.0774
13.9238 | -34.8888 | -23.0857
13.9299 | -34.8893 | -23.0941
13.936 | -34.8898 | -23.1026
13.9421 | -34.8903 | -23.1109
13.9482 | -34.8908 | -23.1193
13.9543 | -34.8913 | -23.1278
13.9604 | -34.8918 | -23.1362
13.9665 | -34.8923 | -23.1445
13.9727 | -34.8928 | -23.1532
13.9788 | -34.8933 | -23.1615
13.9849 | -34.8938 | -23.17
13.991 | -34.8943 | -23.1784
13.9971 | -34.8947 | -23.1867
14.0032 | -34.8952 | -23.1951
14.0093 | -34.8957 | -23.2036
14.0154 | -34.8962 | -23.2119
14.0215 | -34.8967 | -23.2202
14.0276 | -34.8972 | -23.2286
14.0337 | -34.8977 | -23.237
14.0398 | -34.8982 | -23.2455
14.0459 | -34.8987 | -23.2539
14.052 | -34.8992 | -23.2624
14.0581 | -34.8997 | -23.2709
14.0642 | -34.9002 | -23.2794
14.0703 | -34.9007 | -23.2878
14.0764 | -34.9012 | -23.2961
14.0825 | -34.9017 | -23.3045
14.0886 | -34.9022 | -23.313
14.0947 | -34.9027 | -23.3214
14.1008 | -34.9032 | -23.3297
14.1069 | -34.9036 | -23.3379
14.1131 | -34.9041 | -23.3464
14.1192 | -34.9046 | -23.3548
14.1253 | -34.9051 | -23.3631
14.1314 | -34.9056 | -23.3715
14.1375 | -34.9061 | -23.3798
14.1436 | -34.9066 | -23.3882
14.1497 | -34.9071 | -23.3966
14.1558 | -34.9076 | -23.4051
14.1619 | -34.9081 | -23.4135
14.168 | -34.9086 | -23.4218
14.1741 | -34.9091 | -23.4301
14.1802 | -34.9096 | -23.4385
14.1863 | -34.9101 | -23.4469
14.1924 | -34.9106 | -23.4552
14.1985 | -34.9111 | -23.4635
14.2046 | -34.9116 | -23.472
14.2107 | -34.9121 | -23.4801
14.2168 | -34.9126 | -23.4886
14.2229 | -34.9131 | -23.4969
14.229 | -34.9136 | -23.5053
14.2351 | -34.914 | -23.5136
14.2412 | -34.9145 | -23.5221
14.2473 | -34.9151 | -23.5303
14.2534 | -34.9155 | -23.5385
14.2596 | -34.916 | -23.5469
14.2657 | -34.9165 | -23.5552
14.2718 | -34.917 | -23.5636
14.2779 | -34.9175 | -23.572
14.284 | -34.918 | -23.5804
14.2901 | -34.9185 | -23.5887
14.2962 | -34.919 | -23.5972
14.3023 | -34.9195 | -23.6056
14.3084 | -34.92 | -23.614
14.3145 | -34.9205 | -23.6222
14.3206 | -34.9209 | -23.6308
14.3267 | -34.9214 | -23.6392
14.3328 | -34.9219 | -23.6476
14.3389 | -34.9224 | -23.6559
14.345 | -34.9229 | -23.6641
14.3511 | -34.9234 | -23.6725
14.3572 | -34.9239 | -23.681
14.3633 | -34.9244 | -23.6895
14.3694 | -34.9249 | -23.6979
14.3755 | -34.9254 | -23.7063
14.3816 | -34.9259 | -23.7148
14.3877 | -34.9264 | -23.7229
14.3938 | -34.927 | -23.7315
14.4 | -34.9275 | -23.7399
14.4061 | -34.9279 | -23.7482
14.4122 | -34.9284 | -23.7566
14.4183 | -34.9289 | -23.7649
14.4244 | -34.9294 | -23.7733
14.4305 | -34.93 | -23.7816
14.4366 | -34.9305 | -23.79
14.4427 | -34.931 | -23.7985
14.4488 | -34.9315 | -23.8069
14.4549 | -34.932 | -23.8152
14.461 | -34.9325 | -23.8236
14.4671 | -34.933 | -23.832
14.4732 | -34.9335 | -23.8404
14.4793 | -34.934 | -23.8487
14.4854 | -34.9345 | -23.8569
14.4915 | -34.935 | -23.865
14.4976 | -34.9355 | -23.8732
14.5037 | -34.936 | -23.8816
14.5098 | -34.9365 | -23.8899
14.5159 | -34.937 | -23.8982
14.522 | -34.9375 | -23.9064
14.5281 | -34.938 | -23.9146
14.5342 | -34.9385 | -23.923
14.5403 | -34.939 | -23.9312
14.5465 | -34.9395 | -23.9397
14.5526 | -34.94 | -23.9481
14.5587 | -34.9405 | -23.9565
14.5648 | -34.941 | -23.9646
14.5709 | -34.9415 | -23.9728
14.577 | -34.942 | -23.9812
14.5831 | -34.9425 | -23.9896
14.5892 | -34.943 | -23.9978
14.5953 | -34.9435 | -24.0059
14.6014 | -34.944 | -24.0143
14.6075 | -34.9445 | -24.0225
14.6136 | -34.945 | -24.031
14.6197 | -34.9454 | -24.0393
14.6258 | -34.9459 | -24.0476
14.6319 | -34.9464 | -24.0558
14.638 | -34.9469 | -24.0641
14.6441 | -34.9474 | -24.0725
14.6502 | -34.9479 | -24.0808
14.6563 | -34.9484 | -24.0892
14.6624 | -34.9489 | -24.0974
14.6685 | -34.9495 | -24.1056
14.6746 | -34.95 | -24.114
14.6807 | -34.9505 | -24.1222
14.6869 | -34.951 | -24.1304
14.693 | -34.9515 | -24.1387
14.6991 | -34.952 | -24.1469
14.7052 | -34.9526 | -24.155
14.7113 | -34.9531 | -24.1633
14.7174 | -34.9536 | -24.1717
14.7235 | -34.9541 | -24.1798
14.7296 | -34.9546 | -24.188
14.7357 | -34.9551 | -24.1963
14.7418 | -34.9556 | -24.2045
14.7479 | -34.9562 | -24.2128
14.754 | -34.9567 | -24.2212
14.7601 | -34.9572 | -24.2294
14.7662 | -34.9577 | -24.2376
14.7723 | -34.9582 | -24.2457
14.7784 | -34.9587 | -24.2541
14.7845 | -34.9592 | -24.2624
14.7906 | -34.9597 | -24.2707
14.7967 | -34.9602 | -24.2791
14.8028 | -34.9607 | -24.2872
14.8089 | -34.9613 | -24.2955
14.815 | -34.9618 | -24.3037
14.8211 | -34.9623 | -24.312
14.8272 | -34.9628 | -24.3204
14.8334 | -34.9633 | -24.3287
14.8395 | -34.9638 | -24.3369
14.8456 | -34.9643 | -24.3451
14.8517 | -34.9648 | -24.3533
14.8578 | -34.9653 | -24.3616
14.8639 | -34.9658 | -24.3698
14.87 | -34.9664 | -24.378
14.8761 | -34.9669 | -24.3862
14.8822 | -34.9674 | -24.3946
14.8883 | -34.9679 | -24.403
14.8944 | -34.9684 | -24.4111
14.9005 | -34.9689 | -24.4194
14.9066 | -34.9694 | -24.4276
14.9127 | -34.97 | -24.4357
14.9188 | -34.9705 | -24.444
14.9249 | -34.9709 | -24.4522
14.931 | -34.9715 | -24.4605
14.9371 | -34.972 | -24.4687
14.9432 | -34.9725 | -24.477
14.9493 | -34.973 | -24.4851
14.9554 | -34.9736 | -24.4934
14.9615 | -34.9741 | -24.5016
14.9676 | -34.9746 | -24.5099
14.9738 | -34.9751 | -24.5179
14.9799 | -34.9756 | -24.5262
14.986 | -34.9761 | -24.5342
14.9921 | -34.9766 | -24.5425
14.9982 | -34.9771 | -24.5506
15.0043 | -34.9776 | -24.5589
15.0104 | -34.9781 | -24.5673
15.0165 | -34.9787 | -24.5756
15.0226 | -34.9792 | -24.5835
15.0287 | -34.9797 | -24.5918
15.0348 | -34.9802 | -24.6
15.0409 | -34.9807 | -24.6082
15.047 | -34.9812 | -24.6164
15.0531 | -34.9817 | -24.6245
15.0592 | -34.9822 | -24.6327
15.0653 | -34.9827 | -24.641
15.0714 | -34.9832 | -24.6491
15.0775 | -34.9838 | -24.6573
15.0836 | -34.9843 | -24.6654
15.0897 | -34.9848 | -24.6734
15.0958 | -34.9854 | -24.6816
15.1019 | -34.9859 | -24.6899
15.108 | -34.9864 | -24.698
15.1141 | -34.9869 | -24.7062
15.1203 | -34.9874 | -24.7144
15.1264 | -34.9879 | -24.7226
15.1325 | -34.9885 | -24.7308
15.1386 | -34.989 | -24.7389
15.1447 | -34.9895 | -24.7471
15.1508 | -34.99 | -24.7554
15.1569 | -34.9905 | -24.7635
15.163 | -34.9911 | -24.7715
15.1691 | -34.9916 | -24.7797
15.1752 | -34.9921 | -24.7878
15.1813 | -34.9926 | -24.796
15.1874 | -34.9932 | -24.8042
15.1935 | -34.9937 | -24.8122
15.1996 | -34.9942 | -24.8203
15.2057 | -34.9947 | -24.8285
15.2118 | -34.9952 | -24.8367
15.2179 | -34.9958 | -24.8449
15.224 | -34.9963 | -24.8531
15.2301 | -34.9968 | -24.861
15.2362 | -34.9973 | -24.8692
15.2423 | -34.9979 | -24.8775
15.2484 | -34.9984 | -24.8857
15.2545 | -34.9989 | -24.8939
15.2607 | -34.9995 | -24.9019
15.2668 | -35 | -24.9099
15.2729 | -35.0005 | -24.918
15.279 | -35.001 | -24.9263
15.2851 | -35.0016 | -24.9344
15.2912 | -35.0021 | -24.9426
15.2973 | -35.0026 | -24.9509
15.3034 | -35.0032 | -24.9589
15.3095 | -35.0037 | -24.9671
15.3156 | -35.0042 | -24.9752
15.3217 | -35.0047 | -24.9833
15.3278 | -35.0053 | -24.9917
15.3339 | -35.0058 | -24.9996
15.34 | -35.0063 | -25.0077
15.3461 | -35.0069 | -25.0158
15.3522 | -35.0074 | -25.0239
15.3583 | -35.0079 | -25.0323
15.3644 | -35.0085 | -25.0403
15.3705 | -35.009 | -25.0481
15.3766 | -35.0095 | -25.0562
15.3827 | -35.01 | -25.0643
15.3888 | -35.0105 | -25.0725
15.3949 | -35.0111 | -25.0806
15.401 | -35.0116 | -25.089
15.4072 | -35.0121 | -25.097
15.4133 | -35.0126 | -25.1051
15.4194 | -35.0131 | -25.1133
15.4255 | -35.0136 | -25.1215
15.4316 | -35.0142 | -25.1296
15.4377 | -35.0147 | -25.1379
15.4438 | -35.0152 | -25.146
15.4499 | -35.0158 | -25.1541
15.456 | -35.0163 | -25.1622
15.4621 | -35.0169 | -25.1704
15.4682 | -35.0174 | -25.1786
15.4743 | -35.0179 | -25.1865
15.4804 | -35.0184 | -25.1947
15.4865 | -35.0189 | -25.2028
15.4926 | -35.0195 | -25.2111
15.4987 | -35.02 | -25.2191
15.5048 | -35.0206 | -25.2272
15.5109 | -35.0211 | -25.2352
15.517 | -35.0216 | -25.2434
15.5231 | -35.0222 | -25.2515
15.5292 | -35.0227 | -25.2596
15.5353 | -35.0232 | -25.2678
15.5414 | -35.0237 | -25.2759
15.5476 | -35.0243 | -25.2837
15.5537 | -35.0248 | -25.2918
15.5598 | -35.0254 | -25.3001
15.5659 | -35.0259 | -25.3081
15.572 | -35.0264 | -25.3162
15.5781 | -35.0269 | -25.3242
15.5842 | -35.0275 | -25.3323
15.5903 | -35.028 | -25.3402
15.5964 | -35.0286 | -25.3483
15.6025 | -35.0291 | -25.3563
15.6086 | -35.0296 | -25.3644
15.6147 | -35.0302 | -25.3726
15.6208 | -35.0307 | -25.3805
15.6269 | -35.0312 | -25.3885
15.633 | -35.0318 | -25.3965
15.6391 | -35.0323 | -25.4046
15.6452 | -35.0328 | -25.4128
15.6513 | -35.0334 | -25.4206
15.6574 | -35.0339 | -25.4288
15.6635 | -35.0345 | -25.4368
15.6696 | -35.035 | -25.4449
15.6757 | -35.0355 | -25.4532
15.6818 | -35.0361 | -25.4612
15.688 | -35.0366 | -25.4692
15.6941 | -35.0371 | -25.4771
15.7002 | -35.0377 | -25.4852
15.7063 | -35.0382 | -25.4932
15.7124 | -35.0387 | -25.5013
15.7185 | -35.0392 | -25.5092
15.7246 | -35.0398 | -25.5172
15.7307 | -35.0403 | -25.5253
15.7368 | -35.0409 | -25.5333
15.7429 | -35.0414 | -25.5414
15.749 | -35.0419 | -25.5495
15.7551 | -35.0425 | -25.5576
15.7612 | -35.0431 | -25.5656
15.7673 | -35.0436 | -25.5736
15.7734 | -35.0441 | -25.5816
15.7795 | -35.0447 | -25.5897
15.7856 | -35.0452 | -25.5977
15.7917 | -35.0458 | -25.6056
15.7978 | -35.0463 | -25.6136
15.8039 | -35.0468 | -25.6215
15.81 | -35.0474 | -25.6296
15.8161 | -35.0479 | -25.6377
15.8222 | -35.0485 | -25.6457
15.8283 | -35.049 | -25.6535
15.8345 | -35.0496 | -25.6615
15.8406 | -35.0501 | -25.6695
15.8467 | -35.0507 | -25.6775
15.8528 | -35.0512 | -25.6855
15.8589 | -35.0518 | -25.6935
15.865 | -35.0524 | -25.7013
15.8711 | -35.0529 | -25.7094
15.8772 | -35.0535 | -25.7174
15.8833 | -35.054 | -25.7254
15.8894 | -35.0545 | -25.7335
15.8955 | -35.055 | -25.7415
15.9016 | -35.0556 | -25.7494
15.9077 | -35.0561 | -25.7574
15.9138 | -35.0566 | -25.7655
15.9199 | -35.0572 | -25.7734
15.926 | -35.0577 | -25.7815
15.9321 | -35.0583 | -25.7894
15.9382 | -35.0588 | -25.7973
15.9443 | -35.0593 | -25.8054
15.9504 | -35.0599 | -25.8134
15.9565 | -35.0604 | -25.8213
15.9626 | -35.0609 | -25.8295
15.9687 | -35.0615 | -25.8372
15.9749 | -35.062 | -25.845
15.981 | -35.0626 | -25.8532
15.9871 | -35.0631 | -25.8611
15.9932 | -35.0636 | -25.869
15.9993 | -35.0642 | -25.8768
16.0054 | -35.0648 | -25.8844
16.0115 | -35.0653 | -25.8923
16.0176 | -35.0658 | -25.9003
16.0237 | -35.0664 | -25.9085
16.0298 | -35.0669 | -25.9164
16.0359 | -35.0674 | -25.9244
16.042 | -35.068 | -25.9322
16.0481 | -35.0686 | -25.9402
16.0542 | -35.0691 | -25.9482
16.0603 | -35.0696 | -25.9562
16.0664 | -35.0702 | -25.9642
16.0725 | -35.0707 | -25.9719
16.0786 | -35.0712 | -25.9798
16.0847 | -35.0718 | -25.9878
16.0908 | -35.0723 | -25.9957
16.0969 | -35.0729 | -26.0037
16.103 | -35.0734 | -26.0117
16.1091 | -35.074 | -26.0196
16.1152 | -35.0745 | -26.0276
16.1214 | -35.0751 | -26.0355
16.1275 | -35.0756 | -26.0434
16.1336 | -35.0761 | -26.0514
16.1397 | -35.0767 | -26.0594
16.1458 | -35.0772 | -26.0671
16.1519 | -35.0778 | -26.0752
16.158 | -35.0783 | -26.0832
16.1641 | -35.0789 | -26.0911
16.1702 | -35.0794 | -26.0989
16.1763 | -35.08 | -26.1069
16.1824 | -35.0805 | -26.1145
16.1885 | -35.0811 | -26.1223
16.1946 | -35.0816 | -26.1301
16.2007 | -35.0822 | -26.138
16.2068 | -35.0827 | -26.146
16.2129 | -35.0833 | -26.154
16.219 | -35.0838 | -26.1617
16.2251 | -35.0844 | -26.1697
16.2312 | -35.0849 | -26.1776
16.2373 | -35.0855 | -26.1855
16.2434 | -35.086 | -26.1934
16.2495 | -35.0866 | -26.2013
16.2556 | -35.0871 | -26.2095
16.2618 | -35.0877 | -26.2173
16.2679 | -35.0882 | -26.2252
16.274 | -35.0887 | -26.2331
16.2801 | -35.0893 | -26.2409
16.2862 | -35.0898 | -26.2488
16.2923 | -35.0904 | -26.2567
16.2984 | -35.0909 | -26.2647
16.3045 | -35.0914 | -26.2726
16.3106 | -35.092 | -26.2805
16.3167 | -35.0925 | -26.2884
16.3228 | -35.0931 | -26.2961
16.3289 | -35.0936 | -26.3041
16.335 | -35.0941 | -26.312
16.3411 | -35.0946 | -26.32
16.3472 | -35.0952 | -26.3281
16.3533 | -35.0958 | -26.3361
16.3594 | -35.0963 | -26.3437
16.3655 | -35.0968 | -26.3516
16.3716 | -35.0973 | -26.3595
16.3777 | -35.0979 | -26.3674
16.3838 | -35.0984 | -26.3753
16.3899 | -35.099 | -26.383
16.396 | -35.0995 | -26.3909
16.4021 | -35.1001 | -26.3988
16.4083 | -35.1006 | -26.4066
16.4144 | -35.1012 | -26.4145
16.4205 | -35.1018 | -26.4224
16.4266 | -35.1023 | -26.4302
16.4327 | -35.1029 | -26.438
16.4388 | -35.1034 | -26.4459
16.4449 | -35.104 | -26.4538
16.451 | -35.1045 | -26.4617
16.4571 | -35.1051 | -26.4697
16.4632 | -35.1056 | -26.4772
16.4693 | -35.1061 | -26.485
16.4754 | -35.1067 | -26.493
16.4815 | -35.1072 | -26.5009
16.4876 | -35.1077 | -26.5086
16.4937 | -35.1083 | -26.5164
16.4998 | -35.1089 | -26.5241
16.5059 | -35.1094 | -26.5319
16.512 | -35.11 | -26.54
16.5181 | -35.1105 | -26.548
16.5242 | -35.1111 | -26.5558
16.5303 | -35.1116 | -26.5637
16.5364 | -35.1122 | -26.5714
16.5425 | -35.1127 | -26.5793
16.5487 | -35.1133 | -26.5871
16.5548 | -35.1138 | -26.5951
16.5609 | -35.1144 | -26.6029
16.567 | -35.1149 | -26.6105
16.5731 | -35.1155 | -26.6184
16.5792 | -35.116 | -26.6263
16.5853 | -35.1166 | -26.6342
16.5914 | -35.1171 | -26.6421
16.5975 | -35.1177 | -26.65
16.6036 | -35.1182 | -26.6577
16.6097 | -35.1188 | -26.6655
16.6158 | -35.1193 | -26.6733
16.6219 | -35.1199 | -26.6812
16.628 | -35.1204 | -26.689
16.6341 | -35.1209 | -26.6969
16.6402 | -35.1215 | -26.7045
16.6463 | -35.122 | -26.7123
16.6524 | -35.1225 | -26.7201
16.6585 | -35.1231 | -26.728
16.6646 | -35.1237 | -26.7358
16.6707 | -35.1242 | -26.7436
16.6768 | -35.1248 | -26.7514
16.6829 | -35.1253 | -26.7591
16.689 | -35.1258 | -26.767
16.6952 | -35.1264 | -26.7749
16.7013 | -35.1269 | -26.7826
16.7074 | -35.1275 | -26.7903
16.7135 | -35.128 | -26.798
16.7196 | -35.1286 | -26.8057
16.7257 | -35.1291 | -26.8134
16.7318 | -35.1297 | -26.8211
16.7379 | -35.1302 | -26.8289
16.744 | -35.1308 | -26.8364
16.7501 | -35.1313 | -26.8441
16.7562 | -35.1319 | -26.8519
16.7623 | -35.1324 | -26.8596
16.7684 | -35.1329 | -26.8674
16.7745 | -35.1335 | -26.8753
16.7806 | -35.134 | -26.8828
16.7867 | -35.1346 | -26.8905
16.7928 | -35.1351 | -26.8984
16.7989 | -35.1357 | -26.9061
16.805 | -35.1362 | -26.9139
16.8111 | -35.1368 | -26.9217
16.8172 | -35.1373 | -26.9295
16.8233 | -35.1379 | -26.9374
16.8294 | -35.1384 | -26.9452
16.8356 | -35.139 | -26.953
16.8417 | -35.1395 | -26.9608
16.8478 | -35.1401 | -26.9685
16.8539 | -35.1406 | -26.9762
16.86 | -35.1412 | -26.9841
16.8661 | -35.1417 | -26.9918
16.8722 | -35.1422 | -26.9998
16.8783 | -35.1428 | -27.0077
16.8844 | -35.1434 | -27.0154
16.8905 | -35.1439 | -27.0231
16.8966 | -35.1444 | -27.0309
16.9027 | -35.145 | -27.0386
16.9088 | -35.1455 | -27.0465
16.9149 | -35.1461 | -27.0544
16.921 | -35.1466 | -27.062
16.9271 | -35.1472 | -27.0699
16.9332 | -35.1477 | -27.0776
16.9393 | -35.1483 | -27.0854
16.9454 | -35.1488 | -27.093
16.9515 | -35.1494 | -27.1009
16.9576 | -35.1499 | -27.1086
16.9637 | -35.1505 | -27.1163
16.9698 | -35.151 | -27.124
16.9759 | -35.1516 | -27.1317
16.9821 | -35.1522 | -27.1397
16.9882 | -35.1527 | -27.1473
16.9943 | -35.1533 | -27.155
17.0004 | -35.1538 | -27.1627
17.0065 | -35.1544 | -27.1704
17.0126 | -35.1549 | -27.1781
17.0187 | -35.1555 | -27.186
17.0248 | -35.156 | -27.1937
17.0309 | -35.1565 | -27.2014
17.037 | -35.1571 | -27.2091
17.0431 | -35.1576 | -27.2169
17.0492 | -35.1581 | -27.2245
17.0553 | -35.1587 | -27.2322
17.0614 | -35.1592 | -27.24
17.0675 | -35.1598 | -27.2478
17.0736 | -35.1603 | -27.2555
17.0797 | -35.1608 | -27.2632
17.0858 | -35.1614 | -27.271
17.0919 | -35.1619 | -27.2788
17.098 | -35.1625 | -27.2864
17.1041 | -35.163 | -27.2939
17.1102 | -35.1636 | -27.3017
17.1163 | -35.1641 | -27.3095
17.1225 | -35.1647 | -27.3171
17.1286 | -35.1652 | -27.3248
17.1347 | -35.1658 | -27.3325
17.1408 | -35.1663 | -27.3405
17.1469 | -35.1669 | -27.3481
17.153 | -35.1674 | -27.3558
17.1591 | -35.1679 | -27.3636
17.1652 | -35.1685 | -27.371
17.1713 | -35.169 | -27.3787
17.1774 | -35.1696 | -27.3864
17.1835 | -35.1701 | -27.3941
17.1896 | -35.1706 | -27.4017
17.1957 | -35.1712 | -27.4094
17.2018 | -35.1717 | -27.4168
17.2079 | -35.1723 | -27.4245
17.214 | -35.1728 | -27.4323
17.2201 | -35.1734 | -27.4401
17.2262 | -35.1739 | -27.4479
17.2323 | -35.1745 | -27.4555
17.2384 | -35.175 | -27.4632
17.2445 | -35.1755 | -27.4707
17.2506 | -35.1761 | -27.4784
17.2567 | -35.1766 | -27.4861
17.2628 | -35.1772 | -27.4936
17.269 | -35.1777 | -27.5011
17.2751 | -35.1783 | -27.5088
17.2812 | -35.1788 | -27.5166
17.2873 | -35.1793 | -27.5244
17.2934 | -35.1799 | -27.5323
17.2995 | -35.1805 | -27.5401
17.3056 | -35.181 | -27.5475
17.3117 | -35.1815 | -27.5552
17.3178 | -35.1821 | -27.5627
17.3239 | -35.1826 | -27.5704
17.33 | -35.1832 | -27.5779
17.3361 | -35.1837 | -27.5856
17.3422 | -35.1843 | -27.5929
17.3483 | -35.1848 | -27.6005
17.3544 | -35.1854 | -27.6082
17.3605 | -35.1859 | -27.6158
17.3666 | -35.1864 | -27.6234
17.3727 | -35.187 | -27.631
17.3788 | -35.1875 | -27.6385
17.3849 | -35.1881 | -27.6461
17.391 | -35.1886 | -27.6538
17.3971 | -35.1891 | -27.6615
17.4032 | -35.1897 | -27.6691
17.4094 | -35.1902 | -27.6765
17.4155 | -35.1908 | -27.6841
17.4216 | -35.1913 | -27.6917
17.4277 | -35.1919 | -27.6994
17.4338 | -35.1924 | -27.7072
17.4399 | -35.1929 | -27.7148
17.446 | -35.1935 | -27.7223
17.4521 | -35.194 | -27.73
17.4582 | -35.1945 | -27.7374
17.4643 | -35.1951 | -27.7451
17.4704 | -35.1956 | -27.7527
17.4765 | -35.1962 | -27.7604
17.4826 | -35.1967 | -27.768
17.4887 | -35.1972 | -27.7757
17.4948 | -35.1978 | -27.7834
17.5009 | -35.1983 | -27.791
17.507 | -35.1989 | -27.7986
17.5131 | -35.1994 | -27.8064
17.5192 | -35.1999 | -27.8139
17.5253 | -35.2005 | -27.8215
17.5314 | -35.201 | -27.8292
17.5375 | -35.2016 | -27.8368
17.5436 | -35.2021 | -27.8446
17.5497 | -35.2026 | -27.8523
17.5559 | -35.2032 | -27.8598
17.562 | -35.2038 | -27.8673
17.5681 | -35.2043 | -27.8749
17.5742 | -35.2049 | -27.8824
17.5803 | -35.2054 | -27.8899
17.5864 | -35.206 | -27.8974
17.5925 | -35.2065 | -27.905
17.5986 | -35.2071 | -27.9126
17.6047 | -35.2076 | -27.9203
17.6108 | -35.2081 | -27.9279
17.6169 | -35.2087 | -27.9355
17.623 | -35.2092 | -27.943
17.6291 | -35.2098 | -27.9504
17.6352 | -35.2103 | -27.9581
17.6413 | -35.2108 | -27.9657
17.6474 | -35.2114 | -27.9734
17.6535 | -35.2119 | -27.981
17.6596 | -35.2125 | -27.9883
17.6657 | -35.213 | -27.996
17.6718 | -35.2136 | -28.0037
17.6779 | -35.2141 | -28.0112
17.684 | -35.2146 | -28.0188
17.6901 | -35.2152 | -28.0264
17.6963 | -35.2157 | -28.0338
17.7024 | -35.2163 | -28.0414
17.7085 | -35.2168 | -28.0489
17.7146 | -35.2173 | -28.0564
17.7207 | -35.2179 | -28.0639
17.7268 | -35.2184 | -28.0712
17.7329 | -35.219 | -28.0788
17.739 | -35.2195 | -28.0864
17.7451 | -35.2201 | -28.0939
17.7512 | -35.2206 | -28.1013
17.7573 | -35.2211 | -28.109
17.7634 | -35.2217 | -28.1161
17.7695 | -35.2222 | -28.1236
17.7756 | -35.2227 | -28.1313
17.7817 | -35.2233 | -28.139
17.7878 | -35.2238 | -28.1465
17.7939 | -35.2243 | -28.1542
17.8 | -35.2249 | -28.1616
17.8061 | -35.2254 | -28.1694
17.8122 | -35.2259 | -28.1768
17.8183 | -35.2265 | -28.1845
17.8244 | -35.227 | -28.1919
17.8305 | -35.2275 | -28.1995
17.8366 | -35.2281 | -28.2067
17.8428 | -35.2286 | -28.2144
17.8489 | -35.2291 | -28.2218
17.855 | -35.2297 | -28.2292
17.8611 | -35.2302 | -28.2366
17.8672 | -35.2308 | -28.244
17.8733 | -35.2313 | -28.2516
17.8794 | -35.2318 | -28.2591
17.8855 | -35.2324 | -28.2667
17.8916 | -35.2329 | -28.2743
17.8977 | -35.2334 | -28.2819
17.9038 | -35.234 | -28.2892
17.9099 | -35.2345 | -28.2968
17.916 | -35.235 | -28.3044
17.9221 | -35.2356 | -28.312
17.9282 | -35.2361 | -28.3195
17.9343 | -35.2366 | -28.3272
17.9404 | -35.2372 | -28.3346
17.9465 | -35.2377 | -28.3422
17.9526 | -35.2382 | -28.3498
17.9587 | -35.2387 | -28.3573
17.9648 | -35.2393 | -28.3648
17.9709 | -35.2398 | -28.3722
17.977 | -35.2403 | -28.3797
17.9832 | -35.2409 | -28.3872
17.9893 | -35.2414 | -28.3948
17.9954 | -35.2419 | -28.4024
18.0015 | -35.2425 | -28.4099
18.0076 | -35.243 | -28.4172
18.0137 | -35.2435 | -28.4247
18.0198 | -35.244 | -28.432
18.0259 | -35.2446 | -28.4395
18.032 | -35.2451 | -28.447
18.0381 | -35.2456 | -28.4545
18.0442 | -35.2462 | -28.4619
18.0503 | -35.2467 | -28.4695
18.0564 | -35.2472 | -28.4769
18.0625 | -35.2478 | -28.4844
18.0686 | -35.2483 | -28.4919
18.0747 | -35.2488 | -28.4995
18.0808 | -35.2494 | -28.5067
18.0869 | -35.2499 | -28.5143
18.093 | -35.2504 | -28.5217
18.0991 | -35.2509 | -28.5293
18.1052 | -35.2515 | -28.5369
18.1113 | -35.252 | -28.5443
18.1174 | -35.2526 | -28.5516
18.1236 | -35.2531 | -28.5591
18.1297 | -35.2537 | -28.5665
18.1358 | -35.2542 | -28.5741
18.1419 | -35.2548 | -28.5816
18.148 | -35.2553 | -28.5892
18.1541 | -35.2558 | -28.5964
18.1602 | -35.2564 | -28.604
18.1663 | -35.2569 | -28.6114
18.1724 | -35.2574 | -28.6189
18.1785 | -35.2579 | -28.6262
18.1846 | -35.2585 | -28.6336
18.1907 | -35.259 | -28.641
18.1968 | -35.2595 | -28.6485
18.2029 | -35.2601 | -28.656
18.209 | -35.2606 | -28.6636
18.2151 | -35.2611 | -28.6712
18.2212 | -35.2616 | -28.6786
18.2273 | -35.2622 | -28.6862
18.2334 | -35.2627 | -28.6937
18.2395 | -35.2633 | -28.701
18.2456 | -35.2638 | -28.7084
18.2517 | -35.2643 | -28.7159
18.2578 | -35.2649 | -28.7231
18.2639 | -35.2654 | -28.7305
18.2701 | -35.266 | -28.738
18.2762 | -35.2665 | -28.7454
18.2823 | -35.267 | -28.7528
18.2884 | -35.2676 | -28.7602
18.2945 | -35.2681 | -28.7672
18.3006 | -35.2686 | -28.7749
18.3067 | -35.2692 | -28.7822
18.3128 | -35.2697 | -28.7897
18.3189 | -35.2702 | -28.7971
18.325 | -35.2708 | -28.8042
18.3311 | -35.2713 | -28.8117
18.3372 | -35.2718 | -28.8192
18.3433 | -35.2724 | -28.8266
18.3494 | -35.2729 | -28.8342
18.3555 | -35.2734 | -28.8417
18.3616 | -35.274 | -28.849
18.3677 | -35.2745 | -28.8565
18.3738 | -35.275 | -28.864
18.3799 | -35.2756 | -28.8714
18.386 | -35.2761 | -28.8789
18.3921 | -35.2766 | -28.8863
18.3982 | -35.2771 | -28.8935
18.4043 | -35.2776 | -28.9009
18.4105 | -35.2782 | -28.9083
18.4166 | -35.2787 | -28.9157
18.4227 | -35.2792 | -28.9231
18.4288 | -35.2797 | -28.9305
18.4349 | -35.2803 | -28.9377
18.441 | -35.2808 | -28.945
18.4471 | -35.2813 | -28.9523
18.4532 | -35.2818 | -28.9597
18.4593 | -35.2824 | -28.9671
18.4654 | -35.2829 | -28.9743
18.4715 | -35.2834 | -28.9817
18.4776 | -35.2839 | -28.9893
18.4837 | -35.2844 | -28.9967
18.4898 | -35.285 | -29.0042
18.4959 | -35.2855 | -29.0118
18.502 | -35.286 | -29.0192
18.5081 | -35.2865 | -29.0266
18.5142 | -35.2871 | -29.0341
18.5203 | -35.2876 | -29.0415
18.5264 | -35.2881 | -29.0489
18.5325 | -35.2886 | -29.0564
18.5386 | -35.2891 | -29.0637
18.5447 | -35.2897 | -29.071
18.5508 | -35.2902 | -29.0784
18.557 | -35.2907 | -29.0859
18.5631 | -35.2912 | -29.0932
18.5692 | -35.2918 | -29.1006
18.5753 | -35.2923 | -29.1079
18.5814 | -35.2928 | -29.1152
18.5875 | -35.2933 | -29.1225
18.5936 | -35.2939 | -29.1298
18.5997 | -35.2944 | -29.1373
18.6058 | -35.2949 | -29.1444
18.6119 | -35.2954 | -29.1518
18.618 | -35.2959 | -29.1593
18.6241 | -35.2965 | -29.1668
18.6302 | -35.297 | -29.1744
18.6363 | -35.2975 | -29.1818
18.6424 | -35.298 | -29.189
18.6485 | -35.2986 | -29.1964
18.6546 | -35.2991 | -29.2038
18.6607 | -35.2996 | -29.2112
18.6668 | -35.3001 | -29.2186
18.6729 | -35.3006 | -29.226
18.679 | -35.3012 | -29.2331
18.6851 | -35.3017 | -29.2403
18.6912 | -35.3022 | -29.2474
18.6974 | -35.3028 | -29.2546
18.7035 | -35.3033 | -29.2619
18.7096 | -35.3038 | -29.2693
18.7157 | -35.3044 | -29.2765
18.7218 | -35.3049 | -29.2837
18.7279 | -35.3054 | -29.2912
18.734 | -35.3059 | -29.2986
18.7401 | -35.3065 | -29.306
18.7462 | -35.307 | -29.313
18.7523 | -35.3075 | -29.3204
18.7584 | -35.308 | -29.3278
18.7645 | -35.3086 | -29.3352
18.7706 | -35.3091 | -29.3425
18.7767 | -35.3096 | -29.3497
18.7828 | -35.3102 | -29.357
18.7889 | -35.3107 | -29.3644
18.795 | -35.3112 | -29.3719
18.8011 | -35.3117 | -29.3793
18.8072 | -35.3123 | -29.3866
18.8133 | -35.3128 | -29.3939
18.8194 | -35.3133 | -29.401
18.8255 | -35.3138 | -29.4083
18.8316 | -35.3143 | -29.4157
18.8377 | -35.3149 | -29.4229
18.8439 | -35.3154 | -29.4302
18.85 | -35.3159 | -29.4375
18.8561 | -35.3164 | -29.4448
18.8622 | -35.317 | -29.4521
18.8683 | -35.3175 | -29.4593
18.8744 | -35.318 | -29.4666
18.8805 | -35.3185 | -29.474
18.8866 | -35.319 | -29.4813
18.8927 | -35.3195 | -29.4885
18.8988 | -35.3201 | -29.4958
18.9049 | -35.3206 | -29.503
18.911 | -35.3211 | -29.5103
18.9171 | -35.3216 | -29.5176
18.9232 | -35.3221 | -29.5246
18.9293 | -35.3227 | -29.5318
18.9354 | -35.3232 | -29.5392
18.9415 | -35.3237 | -29.5465
18.9476 | -35.3242 | -29.5539
18.9537 | -35.3247 | -29.5611
18.9598 | -35.3253 | -29.5682
18.9659 | -35.3258 | -29.5755
18.972 | -35.3263 | -29.5828
18.9781 | -35.3268 | -29.5901
18.9843 | -35.3273 | -29.5974
18.9904 | -35.3279 | -29.6046
18.9965 | -35.3284 | -29.6115
19.0026 | -35.3289 | -29.6189
19.0087 | -35.3294 | -29.6264
19.0148 | -35.33 | -29.6335
19.0209 | -35.3305 | -29.6408
19.027 | -35.331 | -29.6482
19.0331 | -35.3316 | -29.6551
19.0392 | -35.3321 | -29.6623
19.0453 | -35.3326 | -29.6696
19.0514 | -35.3331 | -29.6767
19.0575 | -35.3336 | -29.6841
19.0636 | -35.3341 | -29.6911
19.0697 | -35.3347 | -29.6983
19.0758 | -35.3352 | -29.7055
19.0819 | -35.3357 | -29.7128
19.088 | -35.3363 | -29.7201
19.0941 | -35.3368 | -29.7273
19.1002 | -35.3373 | -29.7346
19.1063 | -35.3378 | -29.7419
19.1124 | -35.3383 | -29.7493
19.1185 | -35.3388 | -29.7565
19.1246 | -35.3393 | -29.7637
19.1308 | -35.3399 | -29.771
19.1369 | -35.3404 | -29.7779
19.143 | -35.3409 | -29.7851
19.1491 | -35.3414 | -29.7925
19.1552 | -35.3419 | -29.8
19.1613 | -35.3425 | -29.8072
19.1674 | -35.343 | -29.8144
19.1735 | -35.3435 | -29.8214
19.1796 | -35.3441 | -29.8286
19.1857 | -35.3446 | -29.8358
19.1918 | -35.3451 | -29.8431
19.1979 | -35.3457 | -29.8503
19.204 | -35.3462 | -29.8573
19.2101 | -35.3467 | -29.8646
19.2162 | -35.3473 | -29.8719
19.2223 | -35.3478 | -29.8789
19.2284 | -35.3483 | -29.8861
19.2345 | -35.3488 | -29.8933
19.2406 | -35.3493 | -29.9001
19.2467 | -35.3498 | -29.9074
19.2528 | -35.3503 | -29.9147
19.2589 | -35.3509 | -29.9218
19.265 | -35.3514 | -29.9291
19.2712 | -35.3519 | -29.9364
19.2773 | -35.3524 | -29.9433
19.2834 | -35.353 | -29.9506
19.2895 | -35.3535 | -29.9578
19.2956 | -35.354 | -29.9649
19.3017 | -35.3545 | -29.9722
19.3078 | -35.355 | -29.9795
19.3139 | -35.3556 | -29.9864
19.32 | -35.3561 | -29.9934
19.3261 | -35.3566 | -30.0005
19.3322 | -35.3571 | -30.0076
19.3383 | -35.3576 | -30.0147
19.3444 | -35.3581 | -30.0215
19.3505 | -35.3587 | -30.0288
19.3566 | -35.3591 | -30.036
19.3627 | -35.3597 | -30.0433
19.3688 | -35.3602 | -30.0505
19.3749 | -35.3608 | -30.0578
19.381 | -35.3613 | -30.0647
19.3871 | -35.3618 | -30.0718
19.3932 | -35.3623 | -30.079
19.3993 | -35.3628 | -30.0861
19.4054 | -35.3633 | -30.0933
19.4115 | -35.3639 | -30.1004
19.4177 | -35.3644 | -30.1073
19.4238 | -35.3649 | -30.1145
19.4299 | -35.3654 | -30.1216
19.436 | -35.3659 | -30.1289
19.4421 | -35.3664 | -30.136
19.4482 | -35.3669 | -30.1433
19.4543 | -35.3675 | -30.15
19.4604 | -35.368 | -30.157
19.4665 | -35.3685 | -30.1641
19.4726 | -35.369 | -30.1713
19.4787 | -35.3695 | -30.1783
19.4848 | -35.37 | -30.1855
19.4909 | -35.3705 | -30.1924
19.497 | -35.371 | -30.1995
19.5031 | -35.3716 | -30.2066
19.5092 | -35.3721 | -30.2137
19.5153 | -35.3726 | -30.2209
19.5214 | -35.3731 | -30.2277
19.5275 | -35.3736 | -30.2348
19.5336 | -35.3741 | -30.2418
19.5397 | -35.3746 | -30.2488
19.5458 | -35.3751 | -30.2559
19.5519 | -35.3756 | -30.2631
19.5581 | -35.3761 | -30.2701
19.5642 | -35.3766 | -30.2772
19.5703 | -35.3771 | -30.2845
19.5764 | -35.3776 | -30.2917
19.5825 | -35.3782 | -30.2986
19.5886 | -35.3787 | -30.3058
19.5947 | -35.3792 | -30.3127
19.6008 | -35.3797 | -30.3198
19.6069 | -35.3802 | -30.3267
19.613 | -35.3807 | -30.3338
19.6191 | -35.3812 | -30.341
19.6252 | -35.3817 | -30.3481
19.6313 | -35.3822 | -30.355
19.6374 | -35.3827 | -30.3621
19.6435 | -35.3832 | -30.3693
19.6496 | -35.3837 | -30.3764
19.6557 | -35.3842 | -30.3834
19.6618 | -35.3847 | -30.39
19.6679 | -35.3852 | -30.3969
19.674 | -35.3857 | -30.404
19.6801 | -35.3862 | -30.4109
19.6862 | -35.3868 | -30.418
19.6923 | -35.3873 | -30.4251
19.6984 | -35.3878 | -30.432
19.7046 | -35.3883 | -30.439
19.7107 | -35.3888 | -30.4461
19.7168 | -35.3893 | -30.4531
19.7229 | -35.3898 | -30.4601
19.729 | -35.3904 | -30.4673
19.7351 | -35.3909 | -30.4739
19.7412 | -35.3914 | -30.4812
19.7473 | -35.3919 | -30.4882
19.7534 | -35.3924 | -30.4952
19.7595 | -35.3929 | -30.5024
19.7656 | -35.3934 | -30.5094
19.7717 | -35.394 | -30.516
19.7778 | -35.3944 | -30.5231
19.7839 | -35.395 | -30.5302
19.79 | -35.3955 | -30.5372
19.7961 | -35.396 | -30.5441
19.8022 | -35.3965 | -30.5512
19.8083 | -35.3971 | -30.5583
19.8144 | -35.3976 | -30.5655
19.8205 | -35.3981 | -30.5724
19.8266 | -35.3986 | -30.5794
19.8327 | -35.3991 | -30.5863
19.8388 | -35.3996 | -30.5928
19.845 | -35.4001 | -30.6
19.8511 | -35.4006 | -30.6071
19.8572 | -35.4012 | -30.6141
19.8633 | -35.4017 | -30.6212
19.8694 | -35.4021 | -30.6283
19.8755 | -35.4026 | -30.6347
19.8816 | -35.403 | -30.6418
19.8877 | -35.4035 | -30.6489
19.8938 | -35.4041 | -30.6559
19.8999 | -35.4046 | -30.6628
19.906 | -35.4051 | -30.6698
19.9121 | -35.4056 | -30.6769
19.9182 | -35.4061 | -30.684
19.9243 | -35.4066 | -30.6913
19.9304 | -35.4071 | -30.6982
19.9365 | -35.4076 | -30.7051
19.9426 | -35.4081 | -30.7115
19.9487 | -35.4086 | -30.7183
19.9548 | -35.4091 | -30.7256
19.9609 | -35.4096 | -30.7324
19.967 | -35.4102 | -30.7393
19.9731 | -35.4107 | -30.7467
19.9792 | -35.4112 | -30.7533
19.9853 | -35.4117 | -30.7602
19.9915 | -35.4122 | -30.7672
19.9976 | -35.4127 | -30.7742
20.0037 | -35.4132 | -30.7812
20.0098 | -35.4136 | -30.7878
20.0159 | -35.4141 | -30.7945
20.022 | -35.4146 | -30.8015
20.0281 | -35.4151 | -30.8083
20.0342 | -35.4157 | -30.8155
20.0403 | -35.4162 | -30.8222
20.0464 | -35.4167 | -30.8293
20.0525 | -35.4172 | -30.8361
20.0586 | -35.4178 | -30.8435
20.0647 | -35.4183 | -30.8499
20.0708 | -35.4188 | -30.8572
20.0769 | -35.4192 | -30.8639
20.083 | -35.4198 | -30.8709
20.0891 | -35.4203 | -30.8779
20.0952 | -35.4208 | -30.8849
20.1013 | -35.4213 | -30.8914
20.1074 | -35.4218 | -30.8978
20.1135 | -35.4223 | -30.9049
20.1196 | -35.4228 | -30.9116
20.1257 | -35.4233 | -30.918
20.1319 | -35.4238 | -30.9251
20.138 | -35.4243 | -30.9319
20.1441 | -35.4248 | -30.9387
20.1502 | -35.4254 | -30.9456
20.1563 | -35.4259 | -30.952
20.1624 | -35.4263 | -30.959
20.1685 | -35.4268 | -30.9659
20.1746 | -35.4273 | -30.9722
20.1807 | -35.4279 | -30.9785
20.1868 | -35.4285 | -30.9856
20.1929 | -35.429 | -30.9928
20.199 | -35.4295 | -30.9999
20.2051 | -35.4301 | -31.0071
20.2112 | -35.4306 | -31.0138
20.2173 | -35.4311 | -31.0208
20.2234 | -35.4315 | -31.0277
20.2295 | -35.432 | -31.0345
20.2356 | -35.4326 | -31.0421
20.2417 | -35.4331 | -31.0497
20.2478 | -35.4336 | -31.0562
20.2539 | -35.4341 | -31.063
20.26 | -35.4347 | -31.0691
20.2661 | -35.4353 | -31.0759
20.2723 | -35.4358 | -31.0824
20.2784 | -35.4363 | -31.0894
20.2845 | -35.4367 | -31.0961
20.2906 | -35.4373 | -31.1028
20.2967 | -35.4378 | -31.1103
20.3028 | -35.4383 | -31.1169
20.3089 | -35.4389 | -31.1234
20.315 | -35.4394 | -31.1299
20.3211 | -35.4399 | -31.1369
20.3272 | -35.4404 | -31.1436
20.3333 | -35.441 | -31.1494
20.3394 | -35.4415 | -31.1556
20.3455 | -35.4421 | -31.1623
20.3516 | -35.4426 | -31.1688
20.3577 | -35.4431 | -31.1758
20.3638 | -35.4435 | -31.1823
20.3699 | -35.444 | -31.1884
20.376 | -35.4445 | -31.1951
20.3821 | -35.445 | -31.2022
20.3882 | -35.4455 | -31.2087
20.3943 | -35.446 | -31.2156
20.4004 | -35.4465 | -31.2224
20.4065 | -35.447 | -31.2291
20.4126 | -35.4474 | -31.2357
20.4188 | -35.4479 | -31.2423
20.4249 | -35.4484 | -31.2487
20.431 | -35.449 | -31.2555
20.4371 | -35.4494 | -31.2614
20.4432 | -35.4499 | -31.2681
20.4493 | -35.4503 | -31.2749
20.4554 | -35.4508 | -31.2815
20.4615 | -35.4512 | -31.2887
20.4676 | -35.4518 | -31.295
20.4737 | -35.4522 | -31.3013
20.4798 | -35.4528 | -31.3083
20.4859 | -35.4534 | -31.3152
20.492 | -35.4539 | -31.3223
20.4981 | -35.4544 | -31.3285
20.5042 | -35.4549 | -31.3355
20.5103 | -35.4555 | -31.3428
20.5164 | -35.4559 | -31.3498
20.5225 | -35.4564 | -31.3567
20.5286 | -35.4568 | -31.3635
20.5347 | -35.4573 | -31.3704
20.5408 | -35.4578 | -31.3776
20.5469 | -35.4584 | -31.3839
20.553 | -35.4589 | -31.3918
20.5592 | -35.4593 | -31.3985
20.5653 | -35.4598 | -31.4054
20.5714 | -35.4602 | -31.4121
20.5775 | -35.4608 | -31.4188
20.5836 | -35.4612 | -31.4245
20.5897 | -35.4619 | -31.4312
20.5958 | -35.4624 | -31.4379
20.6019 | -35.4629 | -31.4452
20.608 | -35.4634 | -31.4514
20.6141 | -35.464 | -31.4585
20.6202 | -35.4644 | -31.4654
20.6263 | -35.4649 | -31.4727
20.6324 | -35.4654 | -31.4799
20.6385 | -35.4658 | -31.4865
20.6446 | -35.4665 | -31.4926
20.6507 | -35.467 | -31.4989
20.6568 | -35.4675 | -31.5052
20.6629 | -35.4679 | -31.5113
20.669 | -35.4685 | -31.5172
20.6751 | -35.469 | -31.524
20.6812 | -35.4696 | -31.5312
20.6873 | -35.4701 | -31.5384
20.6934 | -35.4705 | -31.5455
20.6995 | -35.4709 | -31.5522
20.7057 | -35.4715 | -31.5597
20.7118 | -35.472 | -31.5659
20.7179 | -35.4725 | -31.5717
20.724 | -35.473 | -31.579
20.7301 | -35.4736 | -31.5845
20.7362 | -35.4741 | -31.5917
20.7423 | -35.4748 | -31.5976
20.7484 | -35.4753 | -31.6044
20.7545 | -35.4759 | -31.6108
20.7606 | -35.4765 | -31.618
20.7667 | -35.4771 | -31.6241
20.7728 | -35.4776 | -31.6318
20.7789 | -35.4781 | -31.6385
20.785 | -35.4788 | -31.6449
20.7911 | -35.4791 | -31.6511
20.7972 | -35.4797 | -31.6579
20.8033 | -35.4802 | -31.6643
20.8094 | -35.4807 | -31.671
20.8155 | -35.4813 | -31.6776
20.8216 | -35.4818 | -31.6827
20.8277 | -35.4822 | -31.6888
20.8338 | -35.4828 | -31.6957
20.8399 | -35.4834 | -31.7028
20.8461 | -35.4839 | -31.7101
20.8522 | -35.4843 | -31.7174
20.8583 | -35.4849 | -31.7238
20.8644 | -35.4854 | -31.7296
20.8705 | -35.486 | -31.7371
20.8766 | -35.4864 | -31.7434
20.8827 | -35.4869 | -31.7509
20.8888 | -35.4874 | -31.7574
20.8949 | -35.488 | -31.7629
20.901 | -35.4885 | -31.7691
20.9071 | -35.489 | -31.7757
20.9132 | -35.4896 | -31.7824
20.9193 | -35.4901 | -31.7885
20.9254 | -35.4907 | -31.7947
20.9315 | -35.4913 | -31.8006
20.9376 | -35.492 | -31.8066
20.9437 | -35.4925 | -31.8137
20.9498 | -35.493 | -31.8214
20.9559 | -35.4936 | -31.8275
20.962 | -35.4941 | -31.834
20.9681 | -35.4945 | -31.8388
20.9742 | -35.495 | -31.845
20.9803 | -35.4954 | -31.8519
20.9864 | -35.496 | -31.8582
20.9926 | -35.4965 | -31.8655
20.9987 | -35.497 | -31.8699
21.0048 | -35.4975 | -31.8759
21.0109 | -35.4979 | -31.8819
21.017 | -35.4984 | -31.8883
21.0231 | -35.4988 | -31.8948
21.0292 | -35.4994 | -31.9017
21.0353 | -35.5 | -31.9088
21.0414 | -35.5006 | -31.9154
21.0475 | -35.5011 | -31.9222
21.0536 | -35.5017 | -31.9287
21.0597 | -35.5022 | -31.9356
21.0658 | -35.5027 | -31.9416
21.0719 | -35.5033 | -31.9469
21.078 | -35.5037 | -31.9548
21.0841 | -35.5044 | -31.9604
21.0902 | -35.5049 | -31.9675
21.0963 | -35.5055 | -31.9744
21.1024 | -35.5059 | -31.9814
21.1085 | -35.5063 | -31.9872
21.1146 | -35.5067 | -31.994
21.1207 | -35.5072 | -31.9996
21.1268 | -35.5076 | -32.0062
21.133 | -35.5082 | -32.0125
21.1391 | -35.5088 | -32.0193
21.1452 | -35.5092 | -32.0255
21.1513 | -35.5096 | -32.0317
21.1574 | -35.5102 | -32.0373
21.1635 | -35.5106 | -32.0442
21.1696 | -35.5111 | -32.0509
21.1757 | -35.5117 | -32.0577
21.1818 | -35.5122 | -32.0637
21.1879 | -35.5127 | -32.07
21.194 | -35.5132 | -32.0771
21.2001 | -35.5136 | -32.0833
21.2062 | -35.5141 | -32.0895
21.2123 | -35.5144 | -32.0962
21.2184 | -35.515 | -32.1026
21.2245 | -35.5155 | -32.1091
21.2306 | -35.516 | -32.1164
21.2367 | -35.5166 | -32.1226
21.2428 | -35.5172 | -32.129
21.2489 | -35.5179 | -32.135
21.255 | -35.5184 | -32.1414
21.2611 | -35.5189 | -32.1473
21.2672 | -35.5194 | -32.1535
21.2733 | -35.5199 | -32.1609
21.2795 | -35.5205 | -32.1667
21.2856 | -35.5211 | -32.1725
21.2917 | -35.5217 | -32.1785
21.2978 | -35.5223 | -32.1849
21.3039 | -35.5227 | -32.1912
21.31 | -35.5233 | -32.1974
21.3161 | -35.5237 | -32.2027
21.3222 | -35.5242 | -32.2093
21.3283 | -35.5247 | -32.2161
21.3344 | -35.5252 | -32.2226
21.3405 | -35.5258 | -32.2283
21.3466 | -35.5263 | -32.2344
21.3527 | -35.5269 | -32.2406
21.3588 | -35.5274 | -32.247
21.3649 | -35.5281 | -32.254
21.371 | -35.5285 | -32.2608
21.3771 | -35.529 | -32.2677
21.3832 | -35.5295 | -32.2745
21.3893 | -35.5299 | -32.2805
21.3954 | -35.5304 | -32.2873
21.4015 | -35.5309 | -32.2931
21.4076 | -35.5314 | -32.2991
21.4137 | -35.5319 | -32.3057
21.4199 | -35.5325 | -32.3122
21.426 | -35.533 | -32.3181
21.4321 | -35.5334 | -32.3242
21.4382 | -35.5339 | -32.3309
21.4443 | -35.5343 | -32.3371
21.4504 | -35.5348 | -32.3434
21.4565 | -35.5353 | -32.3493
21.4626 | -35.5358 | -32.3556
21.4687 | -35.5364 | -32.3615
21.4748 | -35.5369 | -32.3678
21.4809 | -35.5374 | -32.3734
21.487 | -35.5379 | -32.3798
21.4931 | -35.5385 | -32.3858
21.4992 | -35.539 | -32.3918
21.5053 | -35.5395 | -32.3986
21.5114 | -35.54 | -32.4042
21.5175 | -35.5404 | -32.4111
21.5236 | -35.5409 | -32.4177
21.5297 | -35.5415 | -32.4232
21.5358 | -35.542 | -32.4296
21.5419 | -35.5425 | -32.4362
21.548 | -35.5431 | -32.4423
21.5541 | -35.5435 | -32.4491
21.5602 | -35.544 | -32.4553
21.5664 | -35.5446 | -32.4611
21.5725 | -35.5452 | -32.4672
21.5786 | -35.5457 | -32.4726
21.5847 | -35.5461 | -32.4791
21.5908 | -35.5466 | -32.4853
21.5969 | -35.547 | -32.4915
21.603 | -35.5475 | -32.4978
21.6091 | -35.548 | -32.504
21.6152 | -35.5486 | -32.5097
21.6213 | -35.5491 | -32.516
21.6274 | -35.5496 | -32.5226
21.6335 | -35.5502 | -32.5282
21.6396 | -35.5506 | -32.5347
21.6457 | -35.5512 | -32.541
21.6518 | -35.5517 | -32.5475
21.6579 | -35.5522 | -32.5537
21.664 | -35.5527 | -32.5601
21.6701 | -35.5533 | -32.5663
21.6762 | -35.5538 | -32.5725
21.6823 | -35.5543 | -32.579
21.6884 | -35.5549 | -32.5859
21.6945 | -35.5554 | -32.592
21.7006 | -35.556 | -32.5978
21.7068 | -35.5564 | -32.6034
21.7129 | -35.557 | -32.6096
21.719 | -35.5575 | -32.616
21.7251 | -35.558 | -32.6226
21.7312 | -35.5585 | -32.6288
21.7373 | -35.5591 | -32.634
21.7434 | -35.5596 | -32.6403
21.7495 | -35.5602 | -32.6466
21.7556 | -35.5607 | -32.6526
21.7617 | -35.5613 | -32.6592
21.7678 | -35.5618 | -32.6653
21.7739 | -35.5623 | -32.6715
21.78 | -35.5629 | -32.6777
21.7861 | -35.5634 | -32.6842
21.7922 | -35.5639 | -32.6905
21.7983 | -35.5645 | -32.6967
21.8044 | -35.565 | -32.7028
21.8105 | -35.5655 | -32.7084
21.8166 | -35.566 | -32.714
21.8227 | -35.5665 | -32.7205
21.8288 | -35.567 | -32.7272
21.8349 | -35.5675 | -32.7334
21.841 | -35.568 | -32.7397
21.8471 | -35.5685 | -32.7448
21.8533 | -35.569 | -32.7512
21.8594 | -35.5696 | -32.7574
21.8655 | -35.5701 | -32.7637
21.8716 | -35.5706 | -32.7697
21.8777 | -35.5711 | -32.7752
21.8838 | -35.5717 | -32.7811
21.8899 | -35.5722 | -32.7875
21.896 | -35.5727 | -32.7936
21.9021 | -35.5733 | -32.7995
21.9082 | -35.5738 | -32.8059
21.9143 | -35.5744 | -32.8118
21.9204 | -35.5749 | -32.8179
21.9265 | -35.5754 | -32.8238
21.9326 | -35.5758 | -32.8299
21.9387 | -35.5764 | -32.8363
21.9448 | -35.5769 | -32.8421
21.9509 | -35.5774 | -32.8485
21.957 | -35.578 | -32.8547
21.9631 | -35.5784 | -32.8609
21.9692 | -35.5789 | -32.8671
21.9753 | -35.5794 | -32.873
21.9814 | -35.5799 | -32.8793
21.9875 | -35.5804 | -32.8851
21.9937 | -35.581 | -32.8913
21.9998 | -35.5815 | -32.8975
22.0059 | -35.582 | -32.9037
22.012 | -35.5826 | -32.9096
22.0181 | -35.5831 | -32.9158
22.0242 | -35.5836 | -32.9216
22.0303 | -35.5841 | -32.9278
22.0364 | -35.5846 | -32.9338
22.0425 | -35.5851 | -32.94
22.0486 | -35.5855 | -32.9461
22.0547 | -35.5861 | -32.9517
22.0608 | -35.5865 | -32.9576
22.0669 | -35.587 | -32.9639
22.073 | -35.5875 | -32.97
22.0791 | -35.588 | -32.976
22.0852 | -35.5885 | -32.9819
22.0913 | -35.5891 | -32.9878
22.0974 | -35.5896 | -32.9936
22.1035 | -35.5902 | -32.9996
22.1096 | -35.5907 | -33.0057
22.1157 | -35.5913 | -33.0122
22.1218 | -35.5918 | -33.0181
22.1279 | -35.5923 | -33.0241
22.134 | -35.5928 | -33.0297
22.1402 | -35.5934 | -33.0357
22.1463 | -35.5939 | -33.042
22.1524 | -35.5943 | -33.048
22.1585 | -35.5949 | -33.0542
22.1646 | -35.5953 | -33.0602
22.1707 | -35.5959 | -33.0668
22.1768 | -35.5964 | -33.073
22.1829 | -35.5969 | -33.0793
22.189 | -35.5975 | -33.0859
22.1951 | -35.598 | -33.0918
22.2012 | -35.5986 | -33.0976
22.2073 | -35.5991 | -33.1039
22.2134 | -35.5996 | -33.1099
22.2195 | -35.6001 | -33.116
22.2256 | -35.6006 | -33.1218
22.2317 | -35.6011 | -33.1275
22.2378 | -35.6016 | -33.1335
22.2439 | -35.6021 | -33.1394
22.25 | -35.6026 | -33.1452
22.2561 | -35.6031 | -33.1514
22.2622 | -35.6036 | -33.1573
22.2683 | -35.6042 | -33.1633
22.2744 | -35.6047 | -33.1693
22.2806 | -35.6052 | -33.1755
22.2867 | -35.6057 | -33.1813
22.2928 | -35.6062 | -33.1872
22.2989 | -35.6068 | -33.1931
22.305 | -35.6073 | -33.1991
22.3111 | -35.6078 | -33.2051
22.3172 | -35.6083 | -33.2115
22.3233 | -35.6089 | -33.2172
22.3294 | -35.6094 | -33.2233
22.3355 | -35.6099 | -33.2289
22.3416 | -35.6104 | -33.2349
22.3477 | -35.611 | -33.2408
22.3538 | -35.6115 | -33.2468
22.3599 | -35.6121 | -33.2529
22.366 | -35.6125 | -33.2589
22.3721 | -35.613 | -33.2648
22.3782 | -35.6135 | -33.2709
22.3843 | -35.614 | -33.277
22.3904 | -35.6145 | -33.2829
22.3965 | -35.6149 | -33.289
22.4026 | -35.6155 | -33.2951
22.4087 | -35.616 | -33.3012
22.4148 | -35.6165 | -33.3073
22.4209 | -35.617 | -33.3133
22.4271 | -35.6175 | -33.3194
22.4332 | -35.618 | -33.3252
22.4393 | -35.6186 | -33.3311
22.4454 | -35.6191 | -33.3366
22.4515 | -35.6196 | -33.3425
22.4576 | -35.6201 | -33.3485
22.4637 | -35.6207 | -33.3547
22.4698 | -35.6212 | -33.3605
22.4759 | -35.6217 | -33.3664
22.482 | -35.6223 | -33.3725
22.4881 | -35.6228 | -33.3785
22.4942 | -35.6233 | -33.3847
22.5003 | -35.6239 | -33.3907
22.5064 | -35.6244 | -33.3966
22.5125 | -35.6248 | -33.4026
22.5186 | -35.6254 | -33.4086
22.5247 | -35.6258 | -33.4146
22.5308 | -35.6263 | -33.4206
22.5369 | -35.6268 | -33.4266
22.543 | -35.6273 | -33.4327
22.5491 | -35.6278 | -33.4381
22.5552 | -35.6284 | -33.4442
22.5613 | -35.6289 | -33.4505
22.5675 | -35.6294 | -33.4563
22.5736 | -35.6299 | -33.4624
22.5797 | -35.6304 | -33.4686
22.5858 | -35.631 | -33.474
22.5919 | -35.6315 | -33.4799
22.598 | -35.6321 | -33.4857
22.6041 | -35.6325 | -33.4916
22.6102 | -35.6331 | -33.4974
22.6163 | -35.6336 | -33.5035
22.6224 | -35.6341 | -33.5096
22.6285 | -35.6346 | -33.5155
22.6346 | -35.6352 | -33.5217
22.6407 | -35.6357 | -33.5276
22.6468 | -35.6362 | -33.5334
22.6529 | -35.6368 | -33.5391
22.659 | -35.6372 | -33.545
22.6651 | -35.6378 | -33.5508
22.6712 | -35.6383 | -33.5567
22.6773 | -35.6388 | -33.5627
22.6834 | -35.6393 | -33.5687
22.6895 | -35.6399 | -33.5743
22.6956 | -35.6404 | -33.5799
22.7017 | -35.641 | -33.5858
22.7079 | -35.6414 | -33.5916
22.714 | -35.642 | -33.5978
22.7201 | -35.6425 | -33.6038
22.7262 | -35.643 | -33.6095
22.7323 | -35.6435 | -33.6153
22.7384 | -35.644 | -33.6212
22.7445 | -35.6445 | -33.6273
22.7506 | -35.645 | -33.6334
22.7567 | -35.6456 | -33.6392
22.7628 | -35.6461 | -33.6448
22.7689 | -35.6466 | -33.6506
22.775 | -35.6471 | -33.6565
22.7811 | -35.6476 | -33.6624
22.7872 | -35.6481 | -33.6682
22.7933 | -35.6487 | -33.674
22.7994 | -35.6492 | -33.6798
22.8055 | -35.6497 | -33.6859
22.8116 | -35.6502 | -33.6917
22.8177 | -35.6507 | -33.6977
22.8238 | -35.6512 | -33.7034
22.8299 | -35.6517 | -33.7094
22.836 | -35.6522 | -33.7152
22.8421 | -35.6528 | -33.7212
22.8482 | -35.6533 | -33.727
22.8544 | -35.6538 | -33.733
22.8605 | -35.6543 | -33.7389
22.8666 | -35.6548 | -33.7445
22.8727 | -35.6553 | -33.7503
22.8788 | -35.6559 | -33.7561
22.8849 | -35.6564 | -33.7619
22.891 | -35.6569 | -33.7676
22.8971 | -35.6574 | -33.7735
22.9032 | -35.6579 | -33.7792
22.9093 | -35.6585 | -33.7852
22.9154 | -35.659 | -33.791
22.9215 | -35.6595 | -33.7966
22.9276 | -35.66 | -33.8024
22.9337 | -35.6605 | -33.808
22.9398 | -35.661 | -33.8138
22.9459 | -35.6615 | -33.8198
22.952 | -35.662 | -33.8257
22.9581 | -35.6626 | -33.8316
22.9642 | -35.6631 | -33.8375
22.9703 | -35.6636 | -33.8433
22.9764 | -35.6641 | -33.8491
22.9825 | -35.6646 | -33.8549
22.9886 | -35.6651 | -33.861
22.9948 | -35.6656 | -33.867
23.0009 | -35.6661 | -33.873
23.007 | -35.6667 | -33.8788
23.0131 | -35.6672 | -33.8846
23.0192 | -35.6677 | -33.8904
23.0253 | -35.6682 | -33.8963
23.0314 | -35.6687 | -33.9024
23.0375 | -35.6692 | -33.9083
23.0436 | -35.6697 | -33.9141
23.0497 | -35.6702 | -33.92
23.0558 | -35.6708 | -33.926
23.0619 | -35.6713 | -33.9318
23.068 | -35.6718 | -33.9375
23.0741 | -35.6723 | -33.9429
23.0802 | -35.6728 | -33.9488
23.0863 | -35.6733 | -33.9547
23.0924 | -35.6738 | -33.9605
23.0985 | -35.6743 | -33.9662
23.1046 | -35.6749 | -33.9722
23.1107 | -35.6754 | -33.978
23.1168 | -35.6759 | -33.984
23.1229 | -35.6764 | -33.9898
23.129 | -35.6769 | -33.9956
23.1351 | -35.6775 | -34.0012
23.1413 | -35.678 | -34.0071
23.1474 | -35.6785 | -34.0129
23.1535 | -35.6789 | -34.0187
23.1596 | -35.6795 | -34.0247
23.1657 | -35.68 | -34.0306
23.1718 | -35.6805 | -34.0364
23.1779 | -35.681 | -34.0423
23.184 | -35.6815 | -34.0481
23.1901 | -35.682 | -34.0541
23.1962 | -35.6826 | -34.0601
23.2023 | -35.6831 | -34.0659
23.2084 | -35.6836 | -34.0717
23.2145 | -35.6841 | -34.0773
23.2206 | -35.6846 | -34.083
23.2267 | -35.6851 | -34.0889
23.2328 | -35.6856 | -34.0946
23.2389 | -35.6861 | -34.1003
23.245 | -35.6866 | -34.1061
23.2511 | -35.6871 | -34.1119
23.2572 | -35.6876 | -34.1176
23.2633 | -35.6881 | -34.1234
23.2694 | -35.6886 | -34.1292
23.2755 | -35.6891 | -34.1351
23.2817 | -35.6896 | -34.141
23.2878 | -35.6901 | -34.1471
23.2939 | -35.6906 | -34.1528
23.3 | -35.6911 | -34.1587
23.3061 | -35.6916 | -34.1645
23.3122 | -35.6921 | -34.1703
23.3183 | -35.6927 | -34.176
23.3244 | -35.6932 | -34.1816
23.3305 | -35.6937 | -34.1874
23.3366 | -35.6942 | -34.1933
23.3427 | -35.6947 | -34.1989
23.3488 | -35.6952 | -34.2046
23.3549 | -35.6957 | -34.2103
23.361 | -35.6963 | -34.2164
23.3671 | -35.6968 | -34.2224
23.3732 | -35.6973 | -34.228
23.3793 | -35.6978 | -34.2337
23.3854 | -35.6984 | -34.2398
23.3915 | -35.6989 | -34.2455
23.3976 | -35.6995 | -34.2514
23.4037 | -35.7 | -34.2572
23.4098 | -35.7005 | -34.263
23.4159 | -35.701 | -34.2688
23.422 | -35.7016 | -34.2746
23.4282 | -35.702 | -34.2803
23.4343 | -35.7025 | -34.2861
23.4404 | -35.703 | -34.2921
23.4465 | -35.7035 | -34.2979
23.4526 | -35.704 | -34.3038
23.4587 | -35.7045 | -34.3097
23.4648 | -35.705 | -34.3153
23.4709 | -35.7055 | -34.3211
23.477 | -35.706 | -34.3271
23.4831 | -35.7065 | -34.3327
23.4892 | -35.707 | -34.3384
23.4953 | -35.7075 | -34.3443
23.5014 | -35.708 | -34.3503
23.5075 | -35.7085 | -34.3561
23.5136 | -35.709 | -34.3618
23.5197 | -35.7096 | -34.3676
23.5258 | -35.7101 | -34.3734
23.5319 | -35.7106 | -34.3792
23.538 | -35.7111 | -34.385
23.5441 | -35.7116 | -34.3907
23.5502 | -35.7121 | -34.3966
23.5563 | -35.7126 | -34.4026
23.5624 | -35.7131 | -34.4084
23.5686 | -35.7136 | -34.4143
23.5747 | -35.7141 | -34.4199
23.5808 | -35.7146 | -34.4257
23.5869 | -35.7151 | -34.4317
23.593 | -35.7156 | -34.4373
23.5991 | -35.7162 | -34.4433
23.6052 | -35.7166 | -34.4489
23.6113 | -35.7171 | -34.455
23.6174 | -35.7176 | -34.4608
23.6235 | -35.7181 | -34.4668
23.6296 | -35.7186 | -34.4726
23.6357 | -35.7191 | -34.4785
23.6418 | -35.7195 | -34.4843
23.6479 | -35.7201 | -34.4898
23.654 | -35.7206 | -34.4956
23.6601 | -35.7211 | -34.5015
23.6662 | -35.7216 | -34.5074
23.6723 | -35.7221 | -34.5131
23.6784 | -35.7226 | -34.5188
23.6845 | -35.7232 | -34.5246
23.6906 | -35.7237 | -34.5304
23.6967 | -35.7242 | -34.5362
23.7028 | -35.7247 | -34.5423
23.7089 | -35.7252 | -34.5481
23.7151 | -35.7257 | -34.554
23.7212 | -35.7262 | -34.5599
23.7273 | -35.7267 | -34.566
23.7334 | -35.7272 | -34.5719
23.7395 | -35.7277 | -34.5778
23.7456 | -35.7282 | -34.5833
23.7517 | -35.7286 | -34.5895
23.7578 | -35.7291 | -34.5954
23.7639 | -35.7297 | -34.6014
23.77 | -35.7302 | -34.6072
23.7761 | -35.7306 | -34.6132
23.7822 | -35.7311 | -34.6192
23.7883 | -35.7316 | -34.625
23.7944 | -35.7321 | -34.631
23.8005 | -35.7326 | -34.6367
23.8066 | -35.7331 | -34.6424
23.8127 | -35.7336 | -34.648
23.8188 | -35.7342 | -34.6541
23.8249 | -35.7347 | -34.6602
23.831 | -35.7352 | -34.6661
23.8371 | -35.7357 | -34.6719
23.8432 | -35.7362 | -34.6775
23.8493 | -35.7367 | -34.6831
23.8555 | -35.7372 | -34.6885
23.8616 | -35.7376 | -34.6943
23.8677 | -35.7382 | -34.7002
23.8738 | -35.7387 | -34.7059
23.8799 | -35.7392 | -34.7119
23.886 | -35.7397 | -34.7171
23.8921 | -35.7403 | -34.7231
23.8982 | -35.7408 | -34.729
23.9043 | -35.7413 | -34.7352
23.9104 | -35.7418 | -34.741
23.9165 | -35.7423 | -34.7466
23.9226 | -35.7427 | -34.7522
23.9287 | -35.7432 | -34.7583
23.9348 | -35.7437 | -34.7642
23.9409 | -35.7442 | -34.7698
23.947 | -35.7447 | -34.7755
23.9531 | -35.7452 | -34.7812
23.9592 | -35.7457 | -34.7872
23.9653 | -35.7462 | -34.793
23.9714 | -35.7467 | -34.7988
23.9775 | -35.7472 | -34.8045
23.9836 | -35.7477 | -34.8104
23.9897 | -35.7483 | -34.8162
23.9958 | -35.7488 | -34.8219
24.002 | -35.7492 | -34.8278
24.0081 | -35.7497 | -34.8336
24.0142 | -35.7502 | -34.8394
24.0203 | -35.7507 | -34.8453
24.0264 | -35.7512 | -34.8508
24.0325 | -35.7517 | -34.8565
24.0386 | -35.7522 | -34.8625
24.0447 | -35.7526 | -34.8683
24.0508 | -35.753 | -34.8739
24.0569 | -35.7536 | -34.8798
24.063 | -35.754 | -34.8855
24.0691 | -35.7545 | -34.8911
24.0752 | -35.755 | -34.8971
24.0813 | -35.7555 | -34.9029
24.0874 | -35.756 | -34.909
24.0935 | -35.7565 | -34.9147
24.0996 | -35.757 | -34.9201
24.1057 | -35.7574 | -34.9259
24.1118 | -35.7579 | -34.9316
24.1179 | -35.7585 | -34.9376
24.124 | -35.7589 | -34.9433
24.1301 | -35.7594 | -34.9491
24.1362 | -35.7599 | -34.9547
24.1424 | -35.7604 | -34.9606
24.1485 | -35.7609 | -34.9664
24.1546 | -35.7614 | -34.9725
24.1607 | -35.7618 | -34.9781
24.1668 | -35.7624 | -34.9836
24.1729 | -35.7628 | -34.9897
24.179 | -35.7634 | -34.9955
24.1851 | -35.7638 | -35.0012
24.1912 | -35.7644 | -35.0068
24.1973 | -35.7649 | -35.0128
24.2034 | -35.7653 | -35.0186
24.2095 | -35.7658 | -35.0243
24.2156 | -35.7663 | -35.0299
24.2217 | -35.7668 | -35.036
24.2278 | -35.7674 | -35.0418
24.2339 | -35.7679 | -35.0474
24.24 | -35.7683 | -35.0526
24.2461 | -35.7689 | -35.0583
24.2522 | -35.7693 | -35.0641
24.2583 | -35.7698 | -35.0699
24.2644 | -35.7703 | -35.0758
24.2705 | -35.7707 | -35.0819
24.2766 | -35.7712 | -35.0876
24.2827 | -35.7717 | -35.0931
24.2889 | -35.7722 | -35.099
24.295 | -35.7727 | -35.1047
24.3011 | -35.7731 | -35.1106
24.3072 | -35.7736 | -35.1167
24.3133 | -35.7741 | -35.1227
24.3194 | -35.7746 | -35.1284
24.3255 | -35.7751 | -35.1343
24.3316 | -35.7756 | -35.1401
24.3377 | -35.7761 | -35.1457
24.3438 | -35.7765 | -35.1517
24.3499 | -35.777 | -35.1576
24.356 | -35.7776 | -35.1632
24.3621 | -35.7781 | -35.1691
24.3682 | -35.7785 | -35.1748
24.3743 | -35.779 | -35.1807
24.3804 | -35.7795 | -35.1862
24.3865 | -35.78 | -35.1915
24.3926 | -35.7805 | -35.1974
24.3987 | -35.7809 | -35.2031
24.4048 | -35.7815 | -35.2088
24.4109 | -35.7819 | -35.2144
24.417 | -35.7825 | -35.2202
24.4231 | -35.783 | -35.2262
24.4293 | -35.7835 | -35.2321
24.4354 | -35.784 | -35.2377
24.4415 | -35.7845 | -35.2435
24.4476 | -35.7849 | -35.2493
24.4537 | -35.7854 | -35.2553
24.4598 | -35.7859 | -35.2613
24.4659 | -35.7863 | -35.2669
24.472 | -35.7868 | -35.2726
24.4781 | -35.7873 | -35.2785
24.4842 | -35.7878 | -35.2842
24.4903 | -35.7883 | -35.2902
24.4964 | -35.7888 | -35.2958
24.5025 | -35.7892 | -35.3016
24.5086 | -35.7897 | -35.3071
24.5147 | -35.7901 | -35.3133
24.5208 | -35.7906 | -35.319
24.5269 | -35.791 | -35.3249
24.533 | -35.7916 | -35.3307
24.5391 | -35.7921 | -35.3365
24.5452 | -35.7926 | -35.3421
24.5513 | -35.7931 | -35.3479
24.5574 | -35.7936 | -35.3536
24.5635 | -35.7941 | -35.3591
24.5696 | -35.7946 | -35.365
24.5758 | -35.7951 | -35.3707
24.5819 | -35.7956 | -35.3765
24.588 | -35.796 | -35.3823
24.5941 | -35.7965 | -35.3881
24.6002 | -35.797 | -35.3938
24.6063 | -35.7974 | -35.3998
24.6124 | -35.7979 | -35.4054
24.6185 | -35.7983 | -35.4115
24.6246 | -35.7988 | -35.4173
24.6307 | -35.7993 | -35.4232
24.6368 | -35.7997 | -35.4291
24.6429 | -35.8002 | -35.4348
24.649 | -35.8007 | -35.4407
24.6551 | -35.8012 | -35.4464
24.6612 | -35.8016 | -35.4525
24.6673 | -35.8021 | -35.4584
24.6734 | -35.8027 | -35.4642
24.6795 | -35.8031 | -35.4704
24.6856 | -35.8036 | -35.4762
24.6917 | -35.8041 | -35.482
24.6978 | -35.8046 | -35.4876
24.7039 | -35.805 | -35.4936
24.71 | -35.8054 | -35.4994
24.7162 | -35.8059 | -35.5054
24.7223 | -35.8064 | -35.5111
24.7284 | -35.8069 | -35.5166
24.7345 | -35.8073 | -35.5224
24.7406 | -35.8078 | -35.5283
24.7467 | -35.8083 | -35.5343
24.7528 | -35.8088 | -35.54
24.7589 | -35.8093 | -35.5458
24.765 | -35.8098 | -35.5517
24.7711 | -35.8103 | -35.5572
24.7772 | -35.8108 | -35.5631
24.7833 | -35.8113 | -35.5689
24.7894 | -35.8118 | -35.5746
24.7955 | -35.8123 | -35.5805
24.8016 | -35.8128 | -35.5858
24.8077 | -35.8133 | -35.5916
24.8138 | -35.8138 | -35.5975
24.8199 | -35.8143 | -35.6034
24.826 | -35.8148 | -35.6094
24.8321 | -35.8152 | -35.6151
24.8382 | -35.8157 | -35.6207
24.8443 | -35.8162 | -35.6265
24.8504 | -35.8167 | -35.6323
24.8565 | -35.8172 | -35.6384
24.8627 | -35.8177 | -35.6443
24.8688 | -35.8182 | -35.6503
24.8749 | -35.8186 | -35.6563
24.881 | -35.8191 | -35.6621
24.8871 | -35.8196 | -35.668
24.8932 | -35.8201 | -35.6736
24.8993 | -35.8206 | -35.6792
24.9054 | -35.821 | -35.6848
24.9115 | -35.8215 | -35.6908
24.9176 | -35.822 | -35.6966
24.9237 | -35.8225 | -35.7026
24.9298 | -35.823 | -35.7084
24.9359 | -35.8234 | -35.7143
24.942 | -35.8239 | -35.7198
24.9481 | -35.8244 | -35.7258
24.9542 | -35.8248 | -35.7315
24.9603 | -35.8252 | -35.7374
24.9664 | -35.8257 | -35.7434
24.9725 | -35.8262 | -35.7495
24.9786 | -35.8267 | -35.7551
24.9847 | -35.8272 | -35.7609
24.9908 | -35.8276 | -35.7666
24.9969 | -35.8281 | -35.7722
25.0031 | -35.8286 | -35.7779
25.0092 | -35.829 | -35.7832
25.0153 | -35.8295 | -35.789
25.0214 | -35.83 | -35.7949
25.0275 | -35.8305 | -35.8006
25.0336 | -35.831 | -35.8064
25.0397 | -35.8314 | -35.8123
25.0458 | -35.8319 | -35.8181
25.0519 | -35.8324 | -35.824
25.058 | -35.8329 | -35.8297
25.0641 | -35.8333 | -35.8355
25.0702 | -35.8338 | -35.8412
25.0763 | -35.8343 | -35.8469
25.0824 | -35.8348 | -35.8524
25.0885 | -35.8352 | -35.8581
25.0946 | -35.8357 | -35.8638
25.1007 | -35.8362 | -35.8697
25.1068 | -35.8366 | -35.8754
25.1129 | -35.8371 | -35.8814
25.119 | -35.8376 | -35.8871
25.1251 | -35.838 | -35.8925
25.1312 | -35.8386 | -35.8983
25.1373 | -35.839 | -35.904
25.1435 | -35.8395 | -35.9098
25.1496 | -35.8399 | -35.9157
25.1557 | -35.8404 | -35.9214
25.1618 | -35.8409 | -35.9269
25.1679 | -35.8414 | -35.9329
25.174 | -35.8418 | -35.9389
25.1801 | -35.8423 | -35.9448
25.1862 | -35.8428 | -35.9507
25.1923 | -35.8433 | -35.9565
25.1984 | -35.8438 | -35.9625
25.2045 | -35.8443 | -35.9685
25.2106 | -35.8447 | -35.9743
25.2167 | -35.8452 | -35.9802
25.2228 | -35.8456 | -35.9859
25.2289 | -35.8461 | -35.9915
25.235 | -35.8465 | -35.9972
25.2411 | -35.847 | -36.0029
25.2472 | -35.8475 | -36.0088
25.2533 | -35.8479 | -36.0147
25.2594 | -35.8483 | -36.0208
25.2655 | -35.8488 | -36.0265
25.2716 | -35.8493 | -36.0326
25.2777 | -35.8497 | -36.0383
25.2838 | -35.8502 | -36.0443
25.29 | -35.8507 | -36.0502
25.2961 | -35.8511 | -36.0562
25.3022 | -35.8516 | -36.0619
25.3083 | -35.8521 | -36.0679
25.3144 | -35.8526 | -36.0733
25.3205 | -35.8531 | -36.0796
25.3266 | -35.8535 | -36.0855
25.3327 | -35.854 | -36.0912
25.3388 | -35.8545 | -36.0973
25.3449 | -35.8549 | -36.1032
25.351 | -35.8554 | -36.1087
25.3571 | -35.8559 | -36.1148
25.3632 | -35.8563 | -36.1205
25.3693 | -35.8568 | -36.1265
25.3754 | -35.8573 | -36.132
25.3815 | -35.8577 | -36.138
25.3876 | -35.8582 | -36.1439
25.3937 | -35.8586 | -36.1495
25.3998 | -35.8591 | -36.1549
25.4059 | -35.8596 | -36.1609
25.412 | -35.86 | -36.1665
25.4181 | -35.8605 | -36.1722
25.4242 | -35.861 | -36.1781
25.4304 | -35.8615 | -36.184
25.4365 | -35.862 | -36.1899
25.4426 | -35.8624 | -36.1959
25.4487 | -35.8629 | -36.2017
25.4548 | -35.8633 | -36.2075
25.4609 | -35.8638 | -36.2134
25.467 | -35.8643 | -36.2193
25.4731 | -35.8648 | -36.2252
25.4792 | -35.8653 | -36.2309
25.4853 | -35.8658 | -36.2366
25.4914 | -35.8663 | -36.2423
25.4975 | -35.8667 | -36.2478
25.5036 | -35.8672 | -36.2535
25.5097 | -35.8677 | -36.2592
25.5158 | -35.8681 | -36.2651
25.5219 | -35.8685 | -36.2711
25.528 | -35.869 | -36.2768
25.5341 | -35.8695 | -36.2828
25.5402 | -35.8699 | -36.2886
25.5463 | -35.8703 | -36.2945
25.5524 | -35.8708 | -36.3
25.5585 | -35.8712 | -36.3058
25.5646 | -35.8717 | -36.3117
25.5707 | -35.8721 | -36.3175
25.5769 | -35.8726 | -36.3233
25.583 | -35.8731 | -36.3292
25.5891 | -35.8736 | -36.3351
25.5952 | -35.874 | -36.341
25.6013 | -35.8745 | -36.3471
25.6074 | -35.875 | -36.3533
25.6135 | -35.8755 | -36.359
25.6196 | -35.876 | -36.3651
25.6257 | -35.8764 | -36.3708
25.6318 | -35.8768 | -36.3767
25.6379 | -35.8773 | -36.3826
25.644 | -35.8777 | -36.3886
25.6501 | -35.8781 | -36.3946
25.6562 | -35.8786 | -36.4004
25.6623 | -35.879 | -36.4063
25.6684 | -35.8795 | -36.4122
25.6745 | -35.8799 | -36.4179
25.6806 | -35.8804 | -36.424
25.6867 | -35.8809 | -36.4299
25.6928 | -35.8814 | -36.4358
25.6989 | -35.8818 | -36.4416
25.705 | -35.8823 | -36.4475
25.7111 | -35.8827 | -36.4535
25.7173 | -35.8832 | -36.4593
25.7234 | -35.8836 | -36.465
25.7295 | -35.8841 | -36.4709
25.7356 | -35.8846 | -36.4766
25.7417 | -35.885 | -36.4823
25.7478 | -35.8855 | -36.4883
25.7539 | -35.8859 | -36.4939
25.76 | -35.8863 | -36.4998
25.7661 | -35.8868 | -36.5055
25.7722 | -35.8872 | -36.5113
25.7783 | -35.8877 | -36.5168
25.7844 | -35.8881 | -36.5225
25.7905 | -35.8886 | -36.5284
25.7966 | -35.889 | -36.5343
25.8027 | -35.8894 | -36.5403
25.8088 | -35.8898 | -36.5459
25.8149 | -35.8903 | -36.5516
25.821 | -35.8907 | -36.5574
25.8271 | -35.8912 | -36.5631
25.8332 | -35.8916 | -36.569
25.8393 | -35.8921 | -36.575
25.8454 | -35.8926 | -36.5808
25.8515 | -35.893 | -36.5869
25.8576 | -35.8935 | -36.592
25.8638 | -35.894 | -36.598
25.8699 | -35.8944 | -36.604
25.876 | -35.8948 | -36.6096
25.8821 | -35.8953 | -36.6155
25.8882 | -35.8957 | -36.6212
25.8943 | -35.8962 | -36.627
25.9004 | -35.8967 | -36.6329
25.9065 | -35.8972 | -36.6389
25.9126 | -35.8977 | -36.6448
25.9187 | -35.8981 | -36.6506
25.9248 | -35.8986 | -36.6561
25.9309 | -35.8991 | -36.6618
25.937 | -35.8995 | -36.6678
25.9431 | -35.9 | -36.6737
25.9492 | -35.9004 | -36.6794
25.9553 | -35.9009 | -36.6854
25.9614 | -35.9013 | -36.6909
25.9675 | -35.9017 | -36.6969
25.9736 | -35.9022 | -36.7028
25.9797 | -35.9026 | -36.7087
25.9858 | -35.9031 | -36.7148
25.9919 | -35.9035 | -36.7204
25.998 | -35.9039 | -36.7263
26.0042 | -35.9043 | -36.732
26.0103 | -35.9048 | -36.7377
26.0164 | -35.9052 | -36.7435
26.0225 | -35.9057 | -36.7495
26.0286 | -35.9061 | -36.7552
26.0347 | -35.9065 | -36.7611
26.0408 | -35.9069 | -36.7667
26.0469 | -35.9074 | -36.7724
26.053 | -35.9078 | -36.7783
26.0591 | -35.9083 | -36.784
26.0652 | -35.9086 | -36.7893
26.0713 | -35.9091 | -36.7947
26.0774 | -35.9095 | -36.8005
26.0835 | -35.91 | -36.8062
26.0896 | -35.9105 | -36.812
26.0957 | -35.9109 | -36.8184
26.1018 | -35.9113 | -36.8244
26.1079 | -35.9118 | -36.8303
26.114 | -35.9122 | -36.8359
26.1201 | -35.9126 | -36.8417
26.1262 | -35.9131 | -36.8477
26.1323 | -35.9135 | -36.8538
26.1384 | -35.914 | -36.859
26.1445 | -35.9144 | -36.865
26.1507 | -35.9149 | -36.8708
26.1568 | -35.9154 | -36.8764
26.1629 | -35.9158 | -36.8822
26.169 | -35.9163 | -36.8881
26.1751 | -35.9167 | -36.8939
26.1812 | -35.9171 | -36.8995
26.1873 | -35.9176 | -36.9052
26.1934 | -35.9181 | -36.911
26.1995 | -35.9186 | -36.9165
26.2056 | -35.9189 | -36.9223
26.2117 | -35.9194 | -36.928
26.2178 | -35.9199 | -36.9335
26.2239 | -35.9203 | -36.9393
26.23 | -35.9207 | -36.9453
26.2361 | -35.9212 | -36.9513
26.2422 | -35.9216 | -36.9568
26.2483 | -35.9221 | -36.9625
26.2544 | -35.9225 | -36.9683
26.2605 | -35.923 | -36.974
26.2666 | -35.9234 | -36.9799
26.2727 | -35.9239 | -36.9859
26.2788 | -35.9243 | -36.9916
26.2849 | -35.9247 | -36.9973
26.2911 | -35.9252 | -37.003
26.2972 | -35.9257 | -37.0087
26.3033 | -35.9261 | -37.0147
26.3094 | -35.9265 | -37.0202
26.3155 | -35.9269 | -37.0257
26.3216 | -35.9274 | -37.0316
26.3277 | -35.9279 | -37.0372
26.3338 | -35.9284 | -37.0429
26.3399 | -35.9288 | -37.0485
26.346 | -35.9292 | -37.0543
26.3521 | -35.9297 | -37.0603
26.3582 | -35.9301 | -37.0658
26.3643 | -35.9306 | -37.0713
26.3704 | -35.9311 | -37.077
26.3765 | -35.9315 | -37.0827
26.3826 | -35.9319 | -37.0883
26.3887 | -35.9323 | -37.094
26.3948 | -35.9327 | -37.0994
26.4009 | -35.9332 | -37.105
26.407 | -35.9337 | -37.1107
26.4131 | -35.9341 | -37.1166
26.4192 | -35.9346 | -37.1224
26.4253 | -35.9351 | -37.1283
26.4314 | -35.9355 | -37.1342
26.4376 | -35.936 | -37.1401
26.4437 | -35.9364 | -37.146
26.4498 | -35.9368 | -37.1518
26.4559 | -35.9373 | -37.157
26.462 | -35.9377 | -37.1625
26.4681 | -35.9382 | -37.1681
26.4742 | -35.9386 | -37.1737
26.4803 | -35.9391 | -37.1792
26.4864 | -35.9395 | -37.1849
26.4925 | -35.94 | -37.1907
26.4986 | -35.9404 | -37.1965
26.5047 | -35.9408 | -37.2024
26.5108 | -35.9412 | -37.2081
26.5169 | -35.9417 | -37.2135
26.523 | -35.9421 | -37.2193
26.5291 | -35.9425 | -37.2247
26.5352 | -35.943 | -37.2304
26.5413 | -35.9434 | -37.2362
26.5474 | -35.9437 | -37.2421
26.5535 | -35.9442 | -37.2481
26.5596 | -35.9446 | -37.2535
26.5657 | -35.945 | -37.2594
26.5718 | -35.9454 | -37.2649
26.578 | -35.9458 | -37.2707
26.5841 | -35.9463 | -37.2758
26.5902 | -35.9467 | -37.2816
26.5963 | -35.947 | -37.2868
26.6024 | -35.9475 | -37.2923
26.6085 | -35.9479 | -37.2983
26.6146 | -35.9484 | -37.3041
26.6207 | -35.9489 | -37.3099
26.6268 | -35.9493 | -37.3159
26.6329 | -35.9497 | -37.3218
26.639 | -35.9502 | -37.3271
26.6451 | -35.9507 | -37.3331
26.6512 | -35.9512 | -37.339
26.6573 | -35.9517 | -37.3439
26.6634 | -35.952 | -37.3493
26.6695 | -35.9524 | -37.3553
26.6756 | -35.9528 | -37.3613
26.6817 | -35.9531 | -37.3668
26.6878 | -35.9535 | -37.3727
26.6939 | -35.954 | -37.3784
26.7 | -35.9544 | -37.3839
26.7061 | -35.9549 | -37.3895
26.7122 | -35.9553 | -37.3954
26.7183 | -35.9557 | -37.401
26.7245 | -35.9561 | -37.4069
26.7306 | -35.9566 | -37.4125
26.7367 | -35.957 | -37.4174
26.7428 | -35.9575 | -37.4231
26.7489 | -35.9579 | -37.4292
26.755 | -35.9583 | -37.435
26.7611 | -35.9587 | -37.4412
26.7672 | -35.9591 | -37.4471
26.7733 | -35.9594 | -37.4525
26.7794 | -35.9599 | -37.4582
26.7855 | -35.9603 | -37.4638
26.7916 | -35.9608 | -37.4696
26.7977 | -35.9613 | -37.4754
26.8038 | -35.9617 | -37.4805
26.8099 | -35.9621 | -37.4861
26.816 | -35.9625 | -37.4917
26.8221 | -35.963 | -37.4972
26.8282 | -35.9635 | -37.5035
26.8343 | -35.9639 | -37.5092
26.8404 | -35.9643 | -37.5145
26.8465 | -35.9648 | -37.5196
26.8526 | -35.9652 | -37.5258
26.8587 | -35.9657 | -37.5319
26.8649 | -35.966 | -37.5371
26.871 | -35.9664 | -37.543
26.8771 | -35.967 | -37.5485
26.8832 | -35.9674 | -37.5541
26.8893 | -35.9678 | -37.5603
26.8954 | -35.9681 | -37.5663
26.9015 | -35.9685 | -37.5725
26.9076 | -35.9688 | -37.5778
26.9137 | -35.9692 | -37.583
26.9198 | -35.9696 | -37.5883
26.9259 | -35.9701 | -37.5942
26.932 | -35.9705 | -37.5999
26.9381 | -35.9709 | -37.6055
26.9442 | -35.9713 | -37.6112
26.9503 | -35.9717 | -37.6171
26.9564 | -35.9721 | -37.6226
26.9625 | -35.9726 | -37.6279
26.9686 | -35.973 | -37.6335
26.9747 | -35.9734 | -37.6394
26.9808 | -35.9738 | -37.6445
26.9869 | -35.9742 | -37.6499
26.993 | -35.9746 | -37.6553
26.9991 | -35.9751 | -37.6607
27.0052 | -35.9755 | -37.6666
27.0114 | -35.976 | -37.6725
27.0175 | -35.9764 | -37.6781
27.0236 | -35.9768 | -37.6844
27.0297 | -35.9774 | -37.6901
27.0358 | -35.9779 | -37.6961
27.0419 | -35.9783 | -37.7017
27.048 | -35.9788 | -37.7079
27.0541 | -35.9793 | -37.7127
27.0602 | -35.9797 | -37.7181
27.0663 | -35.9801 | -37.7241
27.0724 | -35.9805 | -37.7294
27.0785 | -35.981 | -37.735
27.0846 | -35.9815 | -37.7403
27.0907 | -35.9819 | -37.7459
27.0968 | -35.9824 | -37.7517
27.1029 | -35.9827 | -37.7572
27.109 | -35.9831 | -37.7625
27.1151 | -35.9836 | -37.7681
27.1212 | -35.9839 | -37.7739
27.1273 | -35.9844 | -37.7798
27.1334 | -35.9848 | -37.7854
27.1395 | -35.9853 | -37.7912
27.1456 | -35.9857 | -37.7969
27.1518 | -35.9862 | -37.8025
27.1579 | -35.9866 | -37.8076
27.164 | -35.9871 | -37.8126
27.1701 | -35.9877 | -37.818
27.1762 | -35.9881 | -37.823
27.1823 | -35.9884 | -37.8288
27.1884 | -35.9888 | -37.8348
27.1945 | -35.9892 | -37.8398
27.2006 | -35.9896 | -37.8451
27.2067 | -35.9901 | -37.8503
27.2128 | -35.9905 | -37.8556
27.2189 | -35.9909 | -37.8614
27.225 | -35.9913 | -37.8672
27.2311 | -35.9917 | -37.8727
27.2372 | -35.9922 | -37.8786
27.2433 | -35.9927 | -37.8842
27.2494 | -35.9931 | -37.8901
27.2555 | -35.9935 | -37.8961
27.2616 | -35.9939 | -37.9016
27.2677 | -35.9943 | -37.9072
27.2738 | -35.9947 | -37.9124
27.2799 | -35.9951 | -37.9181
27.286 | -35.9955 | -37.9237
27.2921 | -35.9959 | -37.9294
27.2983 | -35.9964 | -37.9346
27.3044 | -35.9967 | -37.9399
27.3105 | -35.9972 | -37.9459
27.3166 | -35.9977 | -37.9513
27.3227 | -35.9981 | -37.9569
27.3288 | -35.9986 | -37.9623
27.3349 | -35.999 | -37.9674
27.341 | -35.9994 | -37.9729
27.3471 | -35.9998 | -37.9784
27.3532 | -36.0003 | -37.9839
27.3593 | -36.0006 | -37.9897
27.3654 | -36.001 | -37.9956
27.3715 | -36.0014 | -38.0015
27.3776 | -36.0018 | -38.0071
27.3837 | -36.0023 | -38.013
27.3898 | -36.0026 | -38.0189
27.3959 | -36.003 | -38.0244
27.402 | -36.0034 | -38.0295
27.4081 | -36.0039 | -38.0351
27.4142 | -36.0043 | -38.0406
27.4203 | -36.0047 | -38.0461
27.4264 | -36.0052 | -38.0514
27.4325 | -36.0055 | -38.0569
27.4387 | -36.006 | -38.0622
27.4448 | -36.0064 | -38.0676
27.4509 | -36.0068 | -38.0728
27.457 | -36.0072 | -38.0785
27.4631 | -36.0076 | -38.084
27.4692 | -36.0081 | -38.0898
27.4753 | -36.0086 | -38.0947
27.4814 | -36.0091 | -38.1004
27.4875 | -36.0095 | -38.1061
27.4936 | -36.0098 | -38.1117
27.4997 | -36.0103 | -38.1174
27.5058 | -36.0107 | -38.1228
27.5119 | -36.0111 | -38.1278
27.518 | -36.0115 | -38.1332
27.5241 | -36.012 | -38.1385
27.5302 | -36.0124 | -38.1441
27.5363 | -36.0127 | -38.1499
27.5424 | -36.0132 | -38.1547
27.5485 | -36.0136 | -38.1599
27.5546 | -36.014 | -38.1657
27.5607 | -36.0144 | -38.1708
27.5668 | -36.0148 | -38.1765
27.5729 | -36.0152 | -38.182
27.5791 | -36.0156 | -38.1868
27.5852 | -36.016 | -38.1926
27.5913 | -36.0164 | -38.198
27.5974 | -36.0168 | -38.2031
27.6035 | -36.0172 | -38.2084
27.6096 | -36.0176 | -38.2141
27.6157 | -36.0179 | -38.2196
27.6218 | -36.0184 | -38.2255
27.6279 | -36.0187 | -38.2312
27.634 | -36.0192 | -38.2365
27.6401 | -36.0196 | -38.242
27.6462 | -36.02 | -38.2475
27.6523 | -36.0204 | -38.2526
27.6584 | -36.0207 | -38.2579
27.6645 | -36.0212 | -38.2634
27.6706 | -36.0216 | -38.2691
27.6767 | -36.0219 | -38.2744
27.6828 | -36.0223 | -38.2804
27.6889 | -36.0227 | -38.2859
27.695 | -36.0231 | -38.2915
27.7011 | -36.0235 | -38.2968
27.7072 | -36.0239 | -38.302
27.7133 | -36.0243 | -38.3074
27.7194 | -36.0248 | -38.3132
27.7256 | -36.0252 | -38.3189
27.7317 | -36.0256 | -38.3245
27.7378 | -36.026 | -38.3296
27.7439 | -36.0264 | -38.3355
27.75 | -36.0268 | -38.3414
27.7561 | -36.0272 | -38.3466
27.7622 | -36.0276 | -38.3519
27.7683 | -36.028 | -38.357
27.7744 | -36.0284 | -38.3621
27.7805 | -36.0289 | -38.3676
27.7866 | -36.0293 | -38.3735
27.7927 | -36.0297 | -38.3778
27.7988 | -36.0301 | -38.3834
27.8049 | -36.0305 | -38.3885
27.811 | -36.0309 | -38.3944
27.8171 | -36.0314 | -38.4
27.8232 | -36.0318 | -38.4054
27.8293 | -36.0323 | -38.4112
27.8354 | -36.0327 | -38.417
27.8415 | -36.0331 | -38.4226
27.8476 | -36.0334 | -38.428
27.8537 | -36.0339 | -38.4341
27.8598 | -36.0343 | -38.4392
27.866 | -36.0347 | -38.4446
27.8721 | -36.0351 | -38.4504
27.8782 | -36.0355 | -38.4563
27.8843 | -36.0359 | -38.4617
27.8904 | -36.0364 | -38.4671
27.8965 | -36.0369 | -38.4724
27.9026 | -36.0372 | -38.4778
27.9087 | -36.0377 | -38.483
27.9148 | -36.0381 | -38.4876
27.9209 | -36.0385 | -38.4932
27.927 | -36.039 | -38.4986
27.9331 | -36.0394 | -38.5039
27.9392 | -36.0398 | -38.5096
27.9453 | -36.0403 | -38.5153
27.9514 | -36.0407 | -38.5206
27.9575 | -36.041 | -38.5258
27.9636 | -36.0414 | -38.5312
27.9697 | -36.0419 | -38.5366
27.9758 | -36.0423 | -38.542
27.9819 | -36.0426 | -38.5475
27.988 | -36.043 | -38.5531
27.9941 | -36.0434 | -38.5581
28.0002 | -36.0438 | -38.563
28.0063 | -36.0442 | -38.5681
28.0125 | -36.0446 | -38.5732
28.0186 | -36.0451 | -38.5785
28.0247 | -36.0455 | -38.5843
28.0308 | -36.046 | -38.5893
28.0369 | -36.0465 | -38.5945
28.043 | -36.0469 | -38.6002
28.0491 | -36.0472 | -38.6056
28.0552 | -36.0476 | -38.6113
28.0613 | -36.048 | -38.6169
28.0674 | -36.0484 | -38.6221
28.0735 | -36.0488 | -38.6272

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.7152 | -43.6256
21.4908 | -35.6862 | -35.8453
16.4948 | -35.0101 | -28.3132
12.6602 | -34.6127 | -21.9122
9.7171 | -34.376 | -16.8962
7.4582 | -34.2172 | -12.8294
5.7244 | -34.1092 | -9.5346
4.3936 | -34.0357 | -6.8838
3.3722 | -34.0184 | -4.791
2.5883 | -34.018 | -3.2035
1.9866 | -34.0148 | -1.9389
1.5248 | -34.0174 | -0.92261
1.1703 | -34.0227 | -0.25898
1 | -34.0163 | 0.070893
0.89822 | -34.0174 | 0.34785
0.68941 | -34.007 | 0.93854
0.52919 | -34.0028 | 1.3944
0.40612 | -34.0064 | 2.0068
0.31171 | -34.0027 | 2.4592
0.23928 | -34.0129 | 3.1816
0.18358 | -34.0251 | 3.5922
0.14091 | -34.0413 | 4.4694
0.10819 | -34.0712 | 5.6229
0.083017 | -34.0695 | 7.1034
0.063705 | -34.1138 | 8.9739
0.048923 | -34.1734 | 11.4649
0.037479 | -34.292 | 14.6518
0.028801 | -34.4754 | 18.6611
0.022078 | -34.7858 | 23.9861
0.017023 | -35.2592 | 30.3782
0.013018 | -36.0523 | 38.0334
0.010014 | -37.1634 | 45.923

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003663 | -16.7964
0.007326 | -31.3104
0.010989 | -36.1381
0.014651 | -39.9339
0.018314 | -41.9637
0.021977 | -44.6242
0.02564 | -46.5388
0.029303 | -48.0815
0.032966 | -50.0052
0.036628 | -50.792
0.040291 | -52.2863
0.043954 | -52.8712
0.047617 | -54.7776
0.05128 | -55.5259
0.054943 | -57.1023
0.058605 | -57.2414
0.062268 | -56.9219
0.065931 | -57.1189
0.069594 | -57.7495
0.073257 | -58.4874
0.07692 | -59.1472
0.080582 | -59.5364
0.084245 | -60.1878
0.087908 | -61.7625
0.091571 | -60.874
0.095234 | -61.2939
0.098897 | -62.4175
0.10256 | -61.814
0.10622 | -63.0865
0.10988 | -64.0815
0.11355 | -64.6027
0.11721 | -65.179
0.12087 | -65.4847
0.12454 | -65.2843
0.1282 | -66.3282
0.13186 | -66.5377
0.13553 | -66.6773
0.13919 | -67.0303
0.14285 | -66.8484
0.14651 | -66.7901
0.15018 | -67.7878
0.15384 | -68.5167
0.1575 | -69.4583
0.16117 | -69.1089
0.16483 | -68.5993
0.16849 | -68.1827
0.17215 | -68.7173
0.17582 | -69.6169
0.17948 | -69.697
0.18314 | -71.2841
0.1868 | -70.8065
0.19047 | -70.3418
0.19413 | -70.2242
0.19779 | -70.7718
0.20146 | -70.7322
0.20512 | -70.2045
0.20878 | -70.7585
0.21244 | -71.8521
0.21611 | -71.5476
0.21977 | -71.8708
0.22343 | -72.5624
0.2271 | -72.4505
0.23076 | -71.877
0.23442 | -71.4776
0.23808 | -71.7265
0.24175 | -71.6385
0.24541 | -71.6944
0.24907 | -71.8412
0.25274 | -72.4573
0.2564 | -72.522
0.26006 | -73.1362
0.26372 | -73.0791
0.26739 | -72.6473
0.27105 | -72.6365
0.27471 | -73.1206
0.27838 | -74.4296
0.28204 | -74.6947
0.2857 | -74.528
0.28936 | -73.9627
0.29303 | -73.8625
0.29669 | -74.0648
0.30035 | -75.0247
0.30401 | -75.0919
0.30768 | -75.2649
0.31134 | -74.9665
0.315 | -75.6689
0.31867 | -75.5654
0.32233 | -75.3058
0.32599 | -75.187
0.32965 | -75.3464
0.33332 | -75.0915
0.33698 | -75.4996
0.34064 | -75.151
0.34431 | -75.6872
0.34797 | -76.2
0.35163 | -75.9277
0.3553 | -75.7033
0.35896 | -74.8403
0.36262 | -75.2556
0.36628 | -75.6073
0.36995 | -75.4177
0.37361 | -76.763
0.37727 | -76.7036
0.38094 | -76.602
0.3846 | -77.5337
0.38826 | -76.894
0.39192 | -77.518
0.39559 | -77.2253
0.39925 | -76.7798
0.40291 | -77.3665
0.40658 | -77.847
0.41024 | -76.8424
0.4139 | -77.753
0.41756 | -78.6173
0.42123 | -77.9098
0.42489 | -78.0672
0.42855 | -77.3711
0.43222 | -77.3596
0.43588 | -77.0054
0.43954 | -77.4822
0.4432 | -76.7944
0.44687 | -77.5123
0.45053 | -78.2032
0.45419 | -78.129
0.45785 | -78.8048
0.46152 | -78.797
0.46518 | -78.9241
0.46884 | -78.5706
0.47251 | -79.2489
0.47617 | -80.0787
0.47983 | -79.9305
0.48349 | -78.7882
0.48716 | -78.8145
0.49082 | -79.9759
0.49448 | -79.5942
0.49815 | -79.4819
0.50181 | -79.6796
0.50547 | -79.3575
0.50913 | -80.6377
0.5128 | -80.4314
0.51646 | -80.7064
0.52012 | -80.0463
0.52379 | -79.5504
0.52745 | -78.9161
0.53111 | -79.0133
0.53477 | -79.6588
0.53844 | -80.2023
0.5421 | -80.2232
0.54576 | -80.2825
0.54943 | -80.4573
0.55309 | -80.333
0.55675 | -80.2555
0.56041 | -80.1365
0.56408 | -80.0887
0.56774 | -80.3773
0.5714 | -81.0204
0.57507 | -81.5319
0.57873 | -81.5817
0.58239 | -81.9607
0.58605 | -82.2415
0.58972 | -81.7169
0.59338 | -81.0184
0.59704 | -80.8573
0.60071 | -81.3937
0.60437 | -82.4004
0.60803 | -82.8863
0.61169 | -82.889
0.61536 | -82.3353
0.61902 | -81.7709
0.62268 | -81.6423
0.62635 | -81.7214
0.63001 | -81.3789
0.63367 | -81.3639
0.63733 | -82.001
0.641 | -82.3556
0.64466 | -82.5601
0.64832 | -82.5879
0.65199 | -82.2197
0.65565 | -82.0048
0.65931 | -82.0481
0.66297 | -82.0031
0.66664 | -81.9863
0.6703 | -82.1787
0.67396 | -82.5698
0.67763 | -83.3948
0.68129 | -84.0605
0.68495 | -83.732
0.68861 | -83.4763
0.69228 | -83.5637
0.69594 | -83.3053
0.6996 | -83.0355
0.70327 | -82.9953
0.70693 | -82.9989
0.71059 | -83.1058
0.71425 | -83.2912
0.71792 | -83.1812
0.72158 | -82.9687
0.72524 | -82.9077
0.7289 | -83.245
0.73257 | -83.8104
0.73623 | -83.8249
0.73989 | -83.2881
0.74356 | -83.1157
0.74722 | -83.6363
0.75088 | -83.4942
0.75454 | -83.2759
0.75821 | -83.4754
0.76187 | -83.8724
0.76553 | -84.4384
0.7692 | -84.4009
0.77286 | -83.9505
0.77652 | -83.9955
0.78018 | -84.6266
0.78385 | -85.001
0.78751 | -84.5374
0.79117 | -84.1868
0.79483 | -84.5357
0.7985 | -84.7364
0.80216 | -84.5414
0.80582 | -84.3803
0.80949 | -84.2722
0.81315 | -84.0918
0.81681 | -84.1542
0.82047 | -84.5986
0.82414 | -85.0335
0.8278 | -84.9122
0.83146 | -84.5915
0.83513 | -84.292
0.83879 | -84.307
0.84245 | -84.6892
0.84611 | -85.1803
0.84978 | -85.1879
0.85344 | -84.9561
0.8571 | -85.044
0.86077 | -85.2051
0.86443 | -84.9939
0.86809 | -84.8102
0.87175 | -85.0592
0.87542 | -85.3486
0.87908 | -85.8077
0.88274 | -86.1842
0.88641 | -85.8303
0.89007 | -85.724
0.89373 | -85.8483
0.8974 | -85.6818
0.90106 | -85.6531
0.90472 | -85.9157
0.90838 | -85.792
0.91205 | -85.5749
0.91571 | -85.6001
0.91937 | -85.421
0.92304 | -85.4432
0.9267 | -85.7151
0.93036 | -85.8496
0.93402 | -85.9845
0.93769 | -86.0827
0.94135 | -86.0983
0.94501 | -86.3065
0.94868 | -86.5027
0.95234 | -86.3962
0.956 | -86.021
0.95966 | -86.1447
0.96333 | -86.4464
0.96699 | -86.4591
0.97065 | -86.3825
0.97431 | -86.4194
0.97798 | -86.0259
0.98164 | -85.4993
0.9853 | -85.3314
0.98897 | -85.2975
0.99263 | -85.5169
0.99629 | -86.2371
0.99995 | -86.952
1.0036 | -87.0772
1.0073 | -86.6889
1.0109 | -86.4272
1.0146 | -86.6598
1.0183 | -86.7923
1.0219 | -86.7468
1.0256 | -86.8626
1.0293 | -86.7429
1.0329 | -86.7386
1.0366 | -86.8817
1.0402 | -86.5321
1.0439 | -86.4145
1.0476 | -86.782
1.0512 | -86.8566
1.0549 | -86.7541
1.0586 | -86.6951
1.0622 | -86.7132
1.0659 | -86.8852
1.0695 | -87.1183
1.0732 | -87.1678
1.0769 | -87.0492
1.0805 | -86.9761
1.0842 | -86.9079
1.0879 | -86.8327
1.0915 | -86.9035
1.0952 | -87.0665
1.0989 | -87.1319
1.1025 | -87.1877
1.1062 | -87.4189
1.1098 | -87.6534
1.1135 | -87.6506
1.1172 | -87.4923
1.1208 | -87.4062
1.1245 | -87.6198
1.1282 | -88.11
1.1318 | -88.4637
1.1355 | -88.3493
1.1391 | -88.0123
1.1428 | -87.7786
1.1465 | -87.7283
1.1501 | -87.9099
1.1538 | -88.319
1.1575 | -88.74
1.1611 | -88.8962
1.1648 | -88.7071
1.1684 | -88.3564
1.1721 | -88.1192
1.1758 | -88.1801
1.1794 | -88.4843
1.1831 | -88.7523
1.1868 | -88.8174
1.1904 | -88.7328
1.1941 | -88.5608
1.1977 | -88.4065
1.2014 | -88.32
1.2051 | -88.163
1.2087 | -87.9314
1.2124 | -87.8258
1.2161 | -87.9403
1.2197 | -88.1617
1.2234 | -88.328
1.227 | -88.4203
1.2307 | -88.5426
1.2344 | -88.7939
1.238 | -89.0312
1.2417 | -88.8383
1.2454 | -88.2771
1.249 | -87.8882
1.2527 | -87.8534
1.2564 | -87.9458
1.26 | -87.9004
1.2637 | -87.6663
1.2673 | -87.4925
1.271 | -87.6832
1.2747 | -88.2075
1.2783 | -88.7588
1.282 | -89.0536
1.2857 | -89.0605
1.2893 | -88.9377
1.293 | -88.7681
1.2966 | -88.5748
1.3003 | -88.513
1.304 | -88.7476
1.3076 | -89.1667
1.3113 | -89.3396
1.315 | -89.1606
1.3186 | -88.9249
1.3223 | -88.6813
1.3259 | -88.3519
1.3296 | -88.1364
1.3333 | -88.28
1.3369 | -88.7356
1.3406 | -89.1952
1.3443 | -89.4991
1.3479 | -89.7159
1.3516 | -89.7921
1.3552 | -89.6781
1.3589 | -89.6325
1.3626 | -89.7536
1.3662 | -89.7037
1.3699 | -89.3925
1.3736 | -89.1514
1.3772 | -89.128
1.3809 | -89.177
1.3846 | -89.1486
1.3882 | -89.1767
1.3919 | -89.3494
1.3955 | -89.4269
1.3992 | -89.3718
1.4029 | -89.4104
1.4065 | -89.5593
1.4102 | -89.7811
1.4139 | -90.1399
1.4175 | -90.5462
1.4212 | -90.7621
1.4248 | -90.692
1.4285 | -90.4379
1.4322 | -90.1653
1.4358 | -89.9245
1.4395 | -89.6987
1.4432 | -89.6268
1.4468 | -89.7624
1.4505 | -89.9251
1.4541 | -90.0957
1.4578 | -90.4466
1.4615 | -90.9219
1.4651 | -91.2755
1.4688 | -91.3042
1.4725 | -90.9962
1.4761 | -90.6035
1.4798 | -90.3914
1.4834 | -90.3778
1.4871 | -90.3859
1.4908 | -90.2851
1.4944 | -90.2101
1.4981 | -90.3835
1.5018 | -90.6928
1.5054 | -90.7669
1.5091 | -90.5838
1.5128 | -90.4582
1.5164 | -90.5475
1.5201 | -90.6997
1.5237 | -90.6515
1.5274 | -90.3037
1.5311 | -89.8462
1.5347 | -89.6393
1.5384 | -89.866
1.5421 | -90.4029
1.5457 | -90.9571
1.5494 | -91.2359
1.553 | -91.0814
1.5567 | -90.6126
1.5604 | -90.136
1.564 | -89.8234
1.5677 | -89.7472
1.5714 | -89.9912
1.575 | -90.4706
1.5787 | -90.8258
1.5823 | -90.8172
1.586 | -90.9181
1.5897 | -91.0767
1.5933 | -91.3035
1.597 | -91.4647
1.6007 | -91.4606
1.6043 | -91.329
1.608 | -91.174
1.6116 | -91.053
1.6153 | -90.9697
1.619 | -90.9077
1.6226 | -90.8483
1.6263 | -90.8135
1.63 | -90.8462
1.6336 | -90.9381
1.6373 | -91.0643
1.641 | -91.2093
1.6446 | -91.3566
1.6483 | -91.5069
1.6519 | -91.6058
1.6556 | -91.5617
1.6593 | -91.4109
1.6629 | -91.3126
1.6666 | -91.3552
1.6703 | -91.4635
1.6739 | -91.5397
1.6776 | -91.5735
1.6812 | -91.5861
1.6849 | -91.5731
1.6886 | -91.4974
1.6922 | -91.348
1.6959 | -91.1735
1.6996 | -91.0449
1.7032 | -91.0037
1.7069 | -91.0187
1.7105 | -91.0302
1.7142 | -91.0308
1.7179 | -91.064
1.7215 | -91.1394
1.7252 | -91.2252
1.7289 | -91.333
1.7325 | -91.4804
1.7362 | -91.6344
1.7398 | -91.7286
1.7435 | -91.7419
1.7472 | -91.6849
1.7508 | -91.6176
1.7545 | -91.6069
1.7582 | -91.6773
1.7618 | -91.8054
1.7655 | -91.9229
1.7692 | -91.9505
1.7728 | -91.8803
1.7765 | -91.7444
1.7801 | -91.5876
1.7838 | -91.462
1.7875 | -91.4583
1.7911 | -91.6115
1.7948 | -91.869
1.7985 | -92.1053
1.8021 | -92.1824
1.8058 | -92.0876
1.8094 | -91.9694
1.8131 | -91.9828
1.8168 | -92.1677
1.8204 | -92.4296
1.8241 | -92.5951
1.8278 | -92.5633
1.8314 | -92.4174
1.8351 | -92.2778
1.8387 | -92.2187
1.8424 | -92.2834
1.8461 | -92.4538
1.8497 | -92.6792
1.8534 | -92.8689
1.8571 | -92.9487
1.8607 | -92.8557
1.8644 | -92.6101
1.868 | -92.3485
1.8717 | -92.1558
1.8754 | -91.9993
1.879 | -91.8365
1.8827 | -91.6754
1.8864 | -91.6016
1.89 | -91.6279
1.8937 | -91.6865
1.8973 | -91.7123
1.901 | -91.704
1.9047 | -91.697
1.9083 | -91.7667
1.912 | -92.0021
1.9157 | -92.4193
1.9193 | -92.8325
1.923 | -93.0404
1.9267 | -93.0735
1.9303 | -93.0871
1.934 | -93.1357
1.9376 | -93.1715
1.9413 | -93.1563
1.945 | -93.0935
1.9486 | -92.9771
1.9523 | -92.7889
1.956 | -92.5143
1.9596 | -92.2056
1.9633 | -91.9742
1.9669 | -91.9184
1.9706 | -92.0402
1.9743 | -92.2627
1.9779 | -92.5283
1.9816 | -92.806
1.9853 | -93.0097
1.9889 | -93.0502
1.9926 | -92.9842
1.9962 | -92.9731
1.9999 | -93.0552
2.0036 | -93.1364
2.0072 | -93.128
2.0109 | -93.0637
2.0146 | -93.0393
2.0182 | -93.1049
2.0219 | -93.2041
2.0255 | -93.2356
2.0292 | -93.1706
2.0329 | -93.048
2.0365 | -92.8963
2.0402 | -92.7812
2.0439 | -92.7814
2.0475 | -92.8898
2.0512 | -92.9775
2.0549 | -92.9322
2.0585 | -92.7662
2.0622 | -92.5867
2.0658 | -92.4994
2.0695 | -92.5154
2.0732 | -92.5755
2.0768 | -92.6454
2.0805 | -92.7408
2.0842 | -92.8839
2.0878 | -93.0453
2.0915 | -93.1953
2.0951 | -93.3408
2.0988 | -93.4588
2.1025 | -93.4967
2.1061 | -93.3951
2.1098 | -93.2008
2.1135 | -93.082
2.1171 | -93.0015
2.1208 | -92.9775
2.1244 | -93.0112
2.1281 | -93.0862
2.1318 | -93.178
2.1354 | -93.2824
2.1391 | -93.3856
2.1428 | -93.4956
2.1464 | -93.6299
2.1501 | -93.7834
2.1537 | -93.9285
2.1574 | -94.0118
2.1611 | -93.9403
2.1647 | -93.715
2.1684 | -93.4322
2.1721 | -93.2109
2.1757 | -93.1074
2.1794 | -93.1305
2.1831 | -93.222
2.1867 | -93.3306
2.1904 | -93.4364
2.194 | -93.5706
2.1977 | -93.7438
2.2014 | -93.9437
2.205 | -94.1183
2.2087 | -94.1805
2.2124 | -94.1116
2.216 | -93.9636
2.2197 | -93.8172
2.2233 | -93.7422
2.227 | -93.7348
2.2307 | -93.7728
2.2343 | -93.8119
2.238 | -93.8267
2.2417 | -93.7898
2.2453 | -93.7228
2.249 | -93.6882
2.2526 | -93.721
2.2563 | -93.8372
2.26 | -94.0176
2.2636 | -94.2361
2.2673 | -94.4432
2.271 | -94.5766
2.2746 | -94.575
2.2783 | -94.4419
2.2819 | -94.2326
2.2856 | -94.0192
2.2893 | -93.8565
2.2929 | -93.772
2.2966 | -93.7868
2.3003 | -93.8995
2.3039 | -94.0862
2.3076 | -94.2811
2.3112 | -94.3882
2.3149 | -94.3866
2.3186 | -94.2939
2.3222 | -94.1653
2.3259 | -94.0203
2.3296 | -93.8635
2.3332 | -93.703
2.3369 | -93.6081
2.3406 | -93.6639
2.3442 | -93.8987
2.3479 | -94.2614
2.3515 | -94.6154
2.3552 | -94.8068
2.3589 | -94.8142
2.3625 | -94.7249
2.3662 | -94.6607
2.3699 | -94.6917
2.3735 | -94.8142
2.3772 | -94.9614
2.3808 | -95.0637
2.3845 | -95.0562
2.3882 | -94.907
2.3918 | -94.6746
2.3955 | -94.4536
2.3992 | -94.304
2.4028 | -94.2407
2.4065 | -94.2621
2.4101 | -94.3593
2.4138 | -94.4989
2.4175 | -94.6329
2.4211 | -94.7106
2.4248 | -94.7145
2.4285 | -94.6442
2.4321 | -94.509
2.4358 | -94.3594
2.4394 | -94.2509
2.4431 | -94.23
2.4468 | -94.2801
2.4504 | -94.3417
2.4541 | -94.3818
2.4578 | -94.3776
2.4614 | -94.3394
2.4651 | -94.2861
2.4688 | -94.2416
2.4724 | -94.2359
2.4761 | -94.2942
2.4797 | -94.4308
2.4834 | -94.5949
2.4871 | -94.7316
2.4907 | -94.809
2.4944 | -94.8268
2.4981 | -94.8164
2.5017 | -94.8146
2.5054 | -94.8118
2.509 | -94.7665
2.5127 | -94.6593
2.5164 | -94.5248
2.52 | -94.4466
2.5237 | -94.4906
2.5274 | -94.6229
2.531 | -94.7335
2.5347 | -94.7452
2.5383 | -94.6637
2.542 | -94.5487
2.5457 | -94.4784
2.5493 | -94.481
2.553 | -94.5499
2.5567 | -94.6845
2.5603 | -94.8565
2.564 | -94.9797
2.5676 | -94.979
2.5713 | -94.857
2.575 | -94.6939
2.5786 | -94.6126
2.5823 | -94.6466
2.586 | -94.7739
2.5896 | -94.9265
2.5933 | -95.0444
2.597 | -95.1267
2.6006 | -95.1912
2.6043 | -95.2789
2.6079 | -95.397
2.6116 | -95.525
2.6153 | -95.6304
2.6189 | -95.6806
2.6226 | -95.6861
2.6263 | -95.6576
2.6299 | -95.6137
2.6336 | -95.5406
2.6372 | -95.4251
2.6409 | -95.2793
2.6446 | -95.1532
2.6482 | -95.044
2.6519 | -94.9477
2.6556 | -94.867
2.6592 | -94.8078
2.6629 | -94.7825
2.6665 | -94.8027
2.6702 | -94.8663
2.6739 | -94.9665
2.6775 | -95.0886
2.6812 | -95.2061
2.6849 | -95.3136
2.6885 | -95.4051
2.6922 | -95.464
2.6958 | -95.4818
2.6995 | -95.4697
2.7032 | -95.4396
2.7068 | -95.4075
2.7105 | -95.3686
2.7142 | -95.3276
2.7178 | -95.2847
2.7215 | -95.2518
2.7252 | -95.2458
2.7288 | -95.2781
2.7325 | -95.344
2.7361 | -95.4071
2.7398 | -95.4634
2.7435 | -95.4754
2.7471 | -95.444
2.7508 | -95.3751
2.7545 | -95.2981
2.7581 | -95.2392
2.7618 | -95.1962
2.7654 | -95.1649
2.7691 | -95.1178
2.7728 | -95.0382
2.7764 | -94.9337
2.7801 | -94.8154
2.7838 | -94.7288
2.7874 | -94.6728
2.7911 | -94.6739
2.7947 | -94.7324
2.7984 | -94.8451
2.8021 | -94.9977
2.8057 | -95.1695
2.8094 | -95.3454
2.8131 | -95.5219
2.8167 | -95.7013
2.8204 | -95.867
2.824 | -95.9857
2.8277 | -96.0315
2.8314 | -95.9782
2.835 | -95.8688
2.8387 | -95.7467
2.8424 | -95.6645
2.846 | -95.6578
2.8497 | -95.7317
2.8533 | -95.8665
2.857 | -96.0093
2.8607 | -96.0935
2.8643 | -96.0777
2.868 | -95.9623
2.8717 | -95.7958
2.8753 | -95.6433
2.879 | -95.5644
2.8827 | -95.5687
2.8863 | -95.6373
2.89 | -95.7384
2.8936 | -95.837
2.8973 | -95.8965
2.901 | -95.9129
2.9046 | -95.8923
2.9083 | -95.8598
2.912 | -95.8081
2.9156 | -95.7417
2.9193 | -95.6595
2.9229 | -95.5667
2.9266 | -95.4684
2.9303 | -95.3829
2.9339 | -95.3087
2.9376 | -95.2659
2.9413 | -95.27
2.9449 | -95.3031
2.9486 | -95.3376
2.9522 | -95.3561
2.9559 | -95.3677
2.9596 | -95.402
2.9632 | -95.4613
2.9669 | -95.5343
2.9706 | -95.6148
2.9742 | -95.7016
2.9779 | -95.7768
2.9815 | -95.8382
2.9852 | -95.8648
2.9889 | -95.8559
2.9925 | -95.8279
2.9962 | -95.7978
2.9999 | -95.7754
3.0035 | -95.7493
3.0072 | -95.7112
3.0109 | -95.6674
3.0145 | -95.6367
3.0182 | -95.6336
3.0218 | -95.6692
3.0255 | -95.7606
3.0292 | -95.8715
3.0328 | -95.9733
3.0365 | -96.0446
3.0402 | -96.0916
3.0438 | -96.0999
3.0475 | -96.0734
3.0511 | -96.023
3.0548 | -95.9648
3.0585 | -95.9098
3.0621 | -95.8871
3.0658 | -95.914
3.0695 | -95.9738
3.0731 | -96.0393
3.0768 | -96.1115
3.0804 | -96.1789
3.0841 | -96.2494
3.0878 | -96.331
3.0914 | -96.4308
3.0951 | -96.531
3.0988 | -96.6047
3.1024 | -96.6584
3.1061 | -96.6942
3.1097 | -96.7169
3.1134 | -96.7319
3.1171 | -96.7387
3.1207 | -96.7323
3.1244 | -96.7022
3.1281 | -96.6646
3.1317 | -96.6335
3.1354 | -96.5928
3.1391 | -96.5505
3.1427 | -96.4984
3.1464 | -96.4379
3.15 | -96.3694
3.1537 | -96.2995
3.1574 | -96.201
3.161 | -96.0826
3.1647 | -95.9866
3.1684 | -95.989
3.172 | -95.9798
3.1757 | -96.0028
3.1793 | -96.0394
3.183 | -96.0835
3.1867 | -96.1428
3.1903 | -96.2162
3.194 | -96.3091
3.1977 | -96.4033
3.2013 | -96.4886
3.205 | -96.5766
3.2086 | -96.677
3.2123 | -96.7652
3.216 | -96.816
3.2196 | -96.8207
3.2233 | -96.7803
3.227 | -96.7147
3.2306 | -96.6359
3.2343 | -96.5708
3.2379 | -96.5399
3.2416 | -96.5528
3.2453 | -96.6043
3.2489 | -96.6847
3.2526 | -96.7762
3.2563 | -96.8397
3.2599 | -96.8482
3.2636 | -96.8111
3.2673 | -96.7501
3.2709 | -96.6607
3.2746 | -96.567
3.2782 | -96.4919
3.2819 | -96.4387
3.2856 | -96.4143
3.2892 | -96.4191
3.2929 | -96.4511
3.2966 | -96.5111
3.3002 | -96.5851
3.3039 | -96.6649
3.3075 | -96.7363
3.3112 | -96.798
3.3149 | -96.8598
3.3185 | -96.9137
3.3222 | -96.9491
3.3259 | -96.9684
3.3295 | -96.9652
3.3332 | -96.9468
3.3368 | -96.926
3.3405 | -96.9095
3.3442 | -96.9022
3.3478 | -96.8965
3.3515 | -96.8878
3.3552 | -96.8864
3.3588 | -96.9069
3.3625 | -96.9497
3.3661 | -97.0136
3.3698 | -97.1015
3.3735 | -97.2058
3.3771 | -97.3069
3.3808 | -97.3855
3.3845 | -97.4189
3.3881 | -97.4004
3.3918 | -97.3514
3.3954 | -97.2888
3.3991 | -97.2276
3.4028 | -97.1923
3.4064 | -97.1989
3.4101 | -97.2411
3.4138 | -97.3106
3.4174 | -97.4032
3.4211 | -97.4953
3.4248 | -97.5525
3.4284 | -97.5627
3.4321 | -97.5299
3.4357 | -97.4577
3.4394 | -97.3533
3.4431 | -97.2425
3.4467 | -97.1287
3.4504 | -97.0329
3.4541 | -96.9717
3.4577 | -96.9652
3.4614 | -97.0164
3.465 | -97.1147
3.4687 | -97.249
3.4724 | -97.4045
3.476 | -97.5523
3.4797 | -97.6814
3.4834 | -97.791
3.487 | -97.8588
3.4907 | -97.878
3.4943 | -97.8428
3.498 | -97.7534
3.5017 | -97.6459
3.5053 | -97.5265
3.509 | -97.42
3.5127 | -97.3276
3.5163 | -97.2543
3.52 | -97.2037
3.5236 | -97.1688
3.5273 | -97.1527
3.531 | -97.162
3.5346 | -97.1812
3.5383 | -97.2008
3.542 | -97.2322
3.5456 | -97.2648
3.5493 | -97.3009
3.553 | -97.3142
3.5566 | -97.2969
3.5603 | -97.2685
3.5639 | -97.2404
3.5676 | -97.2364
3.5713 | -97.2676
3.5749 | -97.3229
3.5786 | -97.382
3.5823 | -97.4414
3.5859 | -97.4887
3.5896 | -97.5175
3.5932 | -97.51
3.5969 | -97.4695
3.6006 | -97.3913
3.6042 | -97.2685
3.6079 | -97.1357
3.6116 | -97.0041
3.6152 | -96.8831
3.6189 | -96.7871
3.6225 | -96.7237
3.6262 | -96.6858
3.6299 | -96.672
3.6335 | -96.6821
3.6372 | -96.7058
3.6409 | -96.7535
3.6445 | -96.8305
3.6482 | -96.9508
3.6518 | -97.0916
3.6555 | -97.2403
3.6592 | -97.3808
3.6628 | -97.489
3.6665 | -97.5675
3.6702 | -97.6216
3.6738 | -97.6559
3.6775 | -97.6673
3.6812 | -97.661
3.6848 | -97.6475
3.6885 | -97.6377
3.6921 | -97.6481
3.6958 | -97.6456
3.6995 | -97.6462
3.7031 | -97.6354
3.7068 | -97.6091
3.7105 | -97.5846
3.7141 | -97.5753
3.7178 | -97.5785
3.7214 | -97.5951
3.7251 | -97.6247
3.7288 | -97.663
3.7324 | -97.7028
3.7361 | -97.7524
3.7398 | -97.8025
3.7434 | -97.8454
3.7471 | -97.8771
3.7507 | -97.8916
3.7544 | -97.9051
3.7581 | -97.9094
3.7617 | -97.9108
3.7654 | -97.9109
3.7691 | -97.8993
3.7727 | -97.8861
3.7764 | -97.8775
3.78 | -97.8805
3.7837 | -97.9015
3.7874 | -97.949
3.791 | -98.023
3.7947 | -98.1137
3.7984 | -98.2021
3.802 | -98.2793
3.8057 | -98.3375
3.8093 | -98.3841
3.813 | -98.4052
3.8167 | -98.3894
3.8203 | -98.3525
3.824 | -98.2949
3.8277 | -98.2234
3.8313 | -98.1422
3.835 | -98.0543
3.8387 | -97.9786
3.8423 | -97.932
3.846 | -97.9204
3.8496 | -97.9471
3.8533 | -98.0017
3.857 | -98.0706
3.8606 | -98.1325
3.8643 | -98.1762
3.868 | -98.1907
3.8716 | -98.1717
3.8753 | -98.1408
3.8789 | -98.1004
3.8826 | -98.0516
3.8863 | -98.0078
3.8899 | -97.9786
3.8936 | -97.984
3.8973 | -98.035
3.9009 | -98.1329
3.9046 | -98.2678
3.9082 | -98.4312
3.9119 | -98.6026
3.9156 | -98.7724
3.9192 | -98.9091
3.9229 | -98.9963
3.9266 | -99.0068
3.9302 | -98.9417
3.9339 | -98.8306
3.9375 | -98.7025
3.9412 | -98.5844
3.9449 | -98.5033
3.9485 | -98.4665
3.9522 | -98.4791
3.9559 | -98.5129
3.9595 | -98.5468
3.9632 | -98.58
3.9669 | -98.6161
3.9705 | -98.6518
3.9742 | -98.6981
3.9778 | -98.7518
3.9815 | -98.8032
3.9852 | -98.8548
3.9888 | -98.9067
3.9925 | -98.956
3.9962 | -98.9944
3.9998 | -99.0167
4.0035 | -99.0198
4.0071 | -98.9897
4.0108 | -98.9412
4.0145 | -98.8693
4.0181 | -98.795
4.0218 | -98.7327
4.0255 | -98.6803
4.0291 | -98.6469
4.0328 | -98.6293
4.0364 | -98.6183
4.0401 | -98.6211
4.0438 | -98.6392
4.0474 | -98.6626
4.0511 | -98.6962
4.0548 | -98.734
4.0584 | -98.7628
4.0621 | -98.7661
4.0657 | -98.7419
4.0694 | -98.6929
4.0731 | -98.6192
4.0767 | -98.5465
4.0804 | -98.4828
4.0841 | -98.4409
4.0877 | -98.4185
4.0914 | -98.4159
4.0951 | -98.4358
4.0987 | -98.4752
4.1024 | -98.5182
4.106 | -98.5587
4.1097 | -98.5991
4.1134 | -98.6239
4.117 | -98.6392
4.1207 | -98.6571
4.1244 | -98.6783
4.128 | -98.6963
4.1317 | -98.7173
4.1353 | -98.7487
4.139 | -98.7852
4.1427 | -98.8211
4.1463 | -98.8468
4.15 | -98.8595
4.1537 | -98.8621
4.1573 | -98.8603
4.161 | -98.8453
4.1646 | -98.8397
4.1683 | -98.8449
4.172 | -98.8472
4.1756 | -98.8696
4.1793 | -98.9007
4.183 | -98.9342
4.1866 | -98.9761
4.1903 | -99.0203
4.1939 | -99.0411
4.1976 | -99.0407
4.2013 | -99.018
4.2049 | -98.9789
4.2086 | -98.9313
4.2123 | -98.8895
4.2159 | -98.8505
4.2196 | -98.8177
4.2233 | -98.8067
4.2269 | -98.8316
4.2306 | -98.8419
4.2342 | -98.8618
4.2379 | -98.8989
4.2416 | -98.9375
4.2452 | -98.9807
4.2489 | -99.0285
4.2526 | -99.0713
4.2562 | -99.1152
4.2599 | -99.1722
4.2635 | -99.2256
4.2672 | -99.2795
4.2709 | -99.3295
4.2745 | -99.3677
4.2782 | -99.3857
4.2819 | -99.3874
4.2855 | -99.3769
4.2892 | -99.3513
4.2928 | -99.3187
4.2965 | -99.2847
4.3002 | -99.2473
4.3038 | -99.2074
4.3075 | -99.1698
4.3112 | -99.1431
4.3148 | -99.128
4.3185 | -99.1163
4.3221 | -99.1213
4.3258 | -99.1384
4.3295 | -99.1632
4.3331 | -99.1998
4.3368 | -99.2465
4.3405 | -99.2921
4.3441 | -99.3378
4.3478 | -99.3836
4.3514 | -99.4262
4.3551 | -99.4545
4.3588 | -99.4718
4.3624 | -99.4705
4.3661 | -99.4562
4.3698 | -99.431
4.3734 | -99.3907
4.3771 | -99.3469
4.3808 | -99.3021
4.3844 | -99.2541
4.3881 | -99.2162
4.3917 | -99.1798
4.3954 | -99.1458
4.3991 | -99.1186
4.4027 | -99.0958
4.4064 | -99.0743
4.4101 | -99.0578
4.4137 | -99.0408
4.4174 | -99.0193
4.421 | -98.9991
4.4247 | -98.989
4.4284 | -98.9846
4.432 | -98.991
4.4357 | -99.0043
4.4394 | -99.0298
4.443 | -99.0549
4.4467 | -99.0738
4.4503 | -99.0946
4.454 | -99.1117
4.4577 | -99.1272
4.4613 | -99.1453
4.465 | -99.1704
4.4687 | -99.1986
4.4723 | -99.2332
4.476 | -99.2745
4.4796 | -99.3185
4.4833 | -99.3576
4.487 | -99.3861
4.4906 | -99.4038
4.4943 | -99.4118
4.498 | -99.4111
4.5016 | -99.4035
4.5053 | -99.403
4.509 | -99.407
4.5126 | -99.4253
4.5163 | -99.4528
4.5199 | -99.4827
4.5236 | -99.5269
4.5273 | -99.5778
4.5309 | -99.6306
4.5346 | -99.6862
4.5383 | -99.7386
4.5419 | -99.7806
4.5456 | -99.8046
4.5492 | -99.8079
4.5529 | -99.7968
4.5566 | -99.7768
4.5602 | -99.7313
4.5639 | -99.6722
4.5676 | -99.6182
4.5712 | -99.5733
4.5749 | -99.5308
4.5785 | -99.5047
4.5822 | -99.4986
4.5859 | -99.5142
4.5895 | -99.5558
4.5932 | -99.6114
4.5969 | -99.6802
4.6005 | -99.7581
4.6042 | -99.8462
4.6078 | -99.9435
4.6115 | -100.0354
4.6152 | -100.1113
4.6188 | -100.171
4.6225 | -100.2078
4.6262 | -100.2261
4.6298 | -100.2345
4.6335 | -100.2105
4.6372 | -100.1585
4.6408 | -100.087
4.6445 | -100.0038
4.6481 | -99.9067
4.6518 | -99.8036
4.6555 | -99.6957
4.6591 | -99.5999
4.6628 | -99.5339
4.6665 | -99.4941
4.6701 | -99.4908
4.6738 | -99.5142
4.6774 | -99.5686
4.6811 | -99.6493
4.6848 | -99.7362
4.6884 | -99.8127
4.6921 | -99.8608
4.6958 | -99.8851
4.6994 | -99.8929
4.7031 | -99.8893
4.7067 | -99.8677
4.7104 | -99.8363
4.7141 | -99.8086
4.7177 | -99.7844
4.7214 | -99.7711
4.7251 | -99.7763
4.7287 | -99.7971
4.7324 | -99.8297
4.736 | -99.8855
4.7397 | -99.9405
4.7434 | -99.9867
4.747 | -100.0115
4.7507 | -100.0156
4.7544 | -99.9748
4.758 | -99.9474
4.7617 | -99.9164
4.7653 | -99.8794
4.769 | -99.8342
4.7727 | -99.7823
4.7763 | -99.7287
4.78 | -99.6794
4.7837 | -99.6334
4.7873 | -99.5818
4.791 | -99.5459
4.7947 | -99.5284
4.7983 | -99.5241
4.802 | -99.5286
4.8056 | -99.541
4.8093 | -99.5522
4.813 | -99.57
4.8166 | -99.5788
4.8203 | -99.5729
4.824 | -99.5601
4.8276 | -99.5456
4.8313 | -99.524
4.8349 | -99.4932
4.8386 | -99.4565
4.8423 | -99.4306
4.8459 | -99.4097
4.8496 | -99.4057
4.8533 | -99.4211
4.8569 | -99.4525
4.8606 | -99.5038
4.8642 | -99.5745
4.8679 | -99.6489
4.8716 | -99.7307
4.8752 | -99.81
4.8789 | -99.8791
4.8826 | -99.9305
4.8862 | -99.9565
4.8899 | -99.9536
4.8935 | -99.9283
4.8972 | -99.8736
4.9009 | -99.8013
4.9045 | -99.7186
4.9082 | -99.6446
4.9119 | -99.5865
4.9155 | -99.5517
4.9192 | -99.5369
4.9229 | -99.5472
4.9265 | -99.5758
4.9302 | -99.6122
4.9338 | -99.6555
4.9375 | -99.6982
4.9412 | -99.7375
4.9448 | -99.7649
4.9485 | -99.7774
4.9522 | -99.7894
4.9558 | -99.8
4.9595 | -99.8071
4.9631 | -99.821
4.9668 | -99.836
4.9705 | -99.8554
4.9741 | -99.89
4.9778 | -99.9205
4.9815 | -99.9592
4.9851 | -99.999
4.9888 | -100.0396
4.9924 | -100.0744
4.9961 | -100.1132
4.9998 | -100.1496
5.0034 | -100.1847
5.0071 | -100.2308
5.0108 | -100.2807
5.0144 | -100.3299
5.0181 | -100.3809
5.0217 | -100.4357
5.0254 | -100.4865
5.0291 | -100.5272
5.0327 | -100.5668
5.0364 | -100.603
5.0401 | -100.6295
5.0437 | -100.6547
5.0474 | -100.68
5.0511 | -100.6982
5.0547 | -100.7155
5.0584 | -100.7266
5.062 | -100.734
5.0657 | -100.7382
5.0694 | -100.7433
5.073 | -100.7406
5.0767 | -100.7383
5.0804 | -100.7339
5.084 | -100.7226
5.0877 | -100.7016
5.0913 | -100.6739
5.095 | -100.6346
5.0987 | -100.5954
5.1023 | -100.5597
5.106 | -100.5381
5.1097 | -100.5265
5.1133 | -100.5339
5.117 | -100.5608
5.1206 | -100.601
5.1243 | -100.6422
5.128 | -100.678
5.1316 | -100.6998
5.1353 | -100.7071
5.139 | -100.6992
5.1426 | -100.677
5.1463 | -100.6484
5.1499 | -100.6169
5.1536 | -100.5807
5.1573 | -100.5458
5.1609 | -100.5177
5.1646 | -100.5025
5.1683 | -100.4926
5.1719 | -100.486
5.1756 | -100.4819
5.1792 | -100.491
5.1829 | -100.494
5.1866 | -100.4969
5.1902 | -100.4935
5.1939 | -100.4801
5.1976 | -100.4572
5.2012 | -100.4344
5.2049 | -100.4157
5.2086 | -100.4042
5.2122 | -100.3992
5.2159 | -100.4043
5.2195 | -100.4199
5.2232 | -100.4395
5.2269 | -100.4615
5.2305 | -100.478
5.2342 | -100.4791
5.2379 | -100.4696
5.2415 | -100.455
5.2452 | -100.4365
5.2488 | -100.4178
5.2525 | -100.4014
5.2562 | -100.3893
5.2598 | -100.3953
5.2635 | -100.4104
5.2672 | -100.4345
5.2708 | -100.4658
5.2745 | -100.5028
5.2781 | -100.5313
5.2818 | -100.5566
5.2855 | -100.585
5.2891 | -100.6178
5.2928 | -100.6567
5.2965 | -100.6942
5.3001 | -100.7355
5.3038 | -100.7728
5.3074 | -100.807
5.3111 | -100.8321
5.3148 | -100.8466
5.3184 | -100.8392
5.3221 | -100.8127
5.3258 | -100.7711
5.3294 | -100.7181
5.3331 | -100.6563
5.3368 | -100.588
5.3404 | -100.5235
5.3441 | -100.4683
5.3477 | -100.4243
5.3514 | -100.3939
5.3551 | -100.3765
5.3587 | -100.3717
5.3624 | -100.3769
5.3661 | -100.386
5.3697 | -100.3965
5.3734 | -100.4162
5.377 | -100.4456
5.3807 | -100.4871
5.3844 | -100.526
5.388 | -100.5698
5.3917 | -100.6129
5.3954 | -100.6565
5.399 | -100.7005
5.4027 | -100.7467
5.4063 | -100.7865
5.41 | -100.8212
5.4137 | -100.8524
5.4173 | -100.8847
5.421 | -100.9196
5.4247 | -100.9448
5.4283 | -100.9655
5.432 | -100.9786
5.4356 | -100.9896
5.4393 | -100.9997
5.443 | -101.0054
5.4466 | -101.0036
5.4503 | -100.9975
5.454 | -100.9863
5.4576 | -100.9611
5.4613 | -100.9316
5.465 | -100.9118
5.4686 | -100.9032
5.4723 | -100.896
5.4759 | -100.8936
5.4796 | -100.9006
5.4833 | -100.9088
5.4869 | -100.9195
5.4906 | -100.9385
5.4943 | -100.9577
5.4979 | -100.9692
5.5016 | -100.969
5.5052 | -100.9698
5.5089 | -100.9645
5.5126 | -100.9533
5.5162 | -100.9388
5.5199 | -100.9125
5.5236 | -100.8793
5.5272 | -100.8467
5.5309 | -100.8093
5.5345 | -100.7737
5.5382 | -100.7471
5.5419 | -100.7286
5.5455 | -100.7102
5.5492 | -100.6916
5.5529 | -100.689
5.5565 | -100.693
5.5602 | -100.691
5.5638 | -100.6783
5.5675 | -100.6631
5.5712 | -100.6476
5.5748 | -100.6246
5.5785 | -100.5996
5.5822 | -100.571
5.5858 | -100.5453
5.5895 | -100.5167
5.5931 | -100.5041
5.5968 | -100.5024
5.6005 | -100.514
5.6041 | -100.5529
5.6078 | -100.6144
5.6115 | -100.6915
5.6151 | -100.7792
5.6188 | -100.8712
5.6225 | -100.9552
5.6261 | -101.0201
5.6298 | -101.0648
5.6334 | -101.089
5.6371 | -101.0814
5.6408 | -101.0581
5.6444 | -101.0278
5.6481 | -100.9895
5.6518 | -100.9402
5.6554 | -100.8858
5.6591 | -100.843
5.6627 | -100.8104
5.6664 | -100.7871
5.6701 | -100.7732
5.6737 | -100.7645
5.6774 | -100.7581
5.6811 | -100.7561
5.6847 | -100.7573
5.6884 | -100.7464
5.692 | -100.7309
5.6957 | -100.7184
5.6994 | -100.7102
5.703 | -100.6993
5.7067 | -100.6974
5.7104 | -100.7089
5.714 | -100.7323
5.7177 | -100.763
5.7214 | -100.7995
5.725 | -100.8374
5.7287 | -100.8757
5.7323 | -100.9075
5.736 | -100.929
5.7397 | -100.9395
5.7433 | -100.9306
5.747 | -100.9106
5.7507 | -100.8852
5.7543 | -100.8549
5.758 | -100.8248
5.7616 | -100.7942
5.7653 | -100.7623
5.769 | -100.7455
5.7726 | -100.7439
5.7763 | -100.7605
5.78 | -100.7879
5.7836 | -100.835
5.7873 | -100.8979
5.7909 | -100.9622
5.7946 | -101.0253
5.7983 | -101.0841
5.8019 | -101.1285
5.8056 | -101.1594
5.8093 | -101.1326
5.8129 | -101.1287
5.8166 | -101.1032
5.8202 | -101.0671
5.8239 | -101.0206
5.8276 | -100.9653
5.8312 | -100.909
5.8349 | -100.8543
5.8386 | -100.8055
5.8422 | -100.7684
5.8459 | -100.7465
5.8495 | -100.7441
5.8532 | -100.7612
5.8569 | -100.7883
5.8605 | -100.837
5.8642 | -100.8952
5.8679 | -100.962
5.8715 | -101.0282
5.8752 | -101.0981
5.8789 | -101.1667
5.8825 | -101.2237
5.8862 | -101.2684
5.8898 | -101.3038
5.8935 | -101.3308
5.8972 | -101.3469
5.9008 | -101.3562
5.9045 | -101.3643
5.9082 | -101.371
5.9118 | -101.3737
5.9155 | -101.3811
5.9191 | -101.382
5.9228 | -101.3769
5.9265 | -101.3709
5.9301 | -101.3599
5.9338 | -101.3432
5.9375 | -101.3249
5.9411 | -101.2988
5.9448 | -101.2659
5.9484 | -101.2278
5.9521 | -101.1875
5.9558 | -101.1545
5.9594 | -101.1364
5.9631 | -101.1308
5.9668 | -101.1371
5.9704 | -101.1592
5.9741 | -101.2018
5.9777 | -101.2465
5.9814 | -101.2948
5.9851 | -101.3465
5.9887 | -101.3927
5.9924 | -101.4285
5.9961 | -101.4554
5.9997 | -101.4728
6.0034 | -101.4839
6.0071 | -101.4767
6.0107 | -101.4703
6.0144 | -101.4658
6.018 | -101.4542
6.0217 | -101.4444
6.0254 | -101.4316
6.029 | -101.4216
6.0327 | -101.4139
6.0364 | -101.4072
6.04 | -101.406
6.0437 | -101.4041
6.0473 | -101.4007
6.051 | -101.3925
6.0547 | -101.3885
6.0583 | -101.389
6.062 | -101.3892
6.0657 | -101.3903
6.0693 | -101.3984
6.073 | -101.4071
6.0766 | -101.414
6.0803 | -101.4244
6.084 | -101.4327
6.0876 | -101.438
6.0913 | -101.4441
6.095 | -101.444
6.0986 | -101.4479
6.1023 | -101.446
6.1059 | -101.446
6.1096 | -101.4553
6.1133 | -101.4668
6.1169 | -101.4885
6.1206 | -101.5179
6.1243 | -101.5571
6.1279 | -101.6002
6.1316 | -101.6376
6.1353 | -101.6776
6.1389 | -101.7152
6.1426 | -101.749
6.1462 | -101.7803
6.1499 | -101.8057
6.1536 | -101.8294
6.1572 | -101.8476
6.1609 | -101.8698
6.1646 | -101.891
6.1682 | -101.9046
6.1719 | -101.9165
6.1755 | -101.9243
6.1792 | -101.925
6.1829 | -101.925
6.1865 | -101.9259
6.1902 | -101.9191
6.1939 | -101.9076
6.1975 | -101.889
6.2012 | -101.8623
6.2048 | -101.8332
6.2085 | -101.8001
6.2122 | -101.7563
6.2158 | -101.7077
6.2195 | -101.6601
6.2232 | -101.617
6.2268 | -101.5703
6.2305 | -101.5247
6.2341 | -101.4768
6.2378 | -101.4271
6.2415 | -101.3822
6.2451 | -101.3404
6.2488 | -101.3059
6.2525 | -101.2692
6.2561 | -101.2365
6.2598 | -101.2134
6.2634 | -101.196
6.2671 | -101.1824
6.2708 | -101.1726
6.2744 | -101.1699
6.2781 | -101.1751
6.2818 | -101.191
6.2854 | -101.2102
6.2891 | -101.2262
6.2928 | -101.2479
6.2964 | -101.2706
6.3001 | -101.2933
6.3037 | -101.3173
6.3074 | -101.3416
6.3111 | -101.3673
6.3147 | -101.3998
6.3184 | -101.4419
6.3221 | -101.4901
6.3257 | -101.5498
6.3294 | -101.6216
6.333 | -101.6953
6.3367 | -101.7667
6.3404 | -101.8049
6.344 | -101.8419
6.3477 | -101.8681
6.3514 | -101.878
6.355 | -101.8751
6.3587 | -101.8578
6.3623 | -101.8352
6.366 | -101.8059
6.3697 | -101.7691
6.3733 | -101.7384
6.377 | -101.7119
6.3807 | -101.703
6.3843 | -101.7094
6.388 | -101.7169
6.3916 | -101.7339
6.3953 | -101.7629
6.399 | -101.8026
6.4026 | -101.8463
6.4063 | -101.8891
6.41 | -101.9259
6.4136 | -101.9612
6.4173 | -101.9978
6.421 | -102.028
6.4246 | -102.0481
6.4283 | -102.0655
6.4319 | -102.0746
6.4356 | -102.0817
6.4393 | -102.0867
6.4429 | -102.0922
6.4466 | -102.0963
6.4503 | -102.0999
6.4539 | -102.103
6.4576 | -102.1086
6.4612 | -102.1169
6.4649 | -102.1304
6.4686 | -102.1478
6.4722 | -102.1657
6.4759 | -102.1879
6.4796 | -102.2057
6.4832 | -102.2221
6.4869 | -102.2332
6.4905 | -102.2448
6.4942 | -102.2581
6.4979 | -102.2638
6.5015 | -102.2719
6.5052 | -102.2851
6.5089 | -102.2995
6.5125 | -102.3108
6.5162 | -102.3226
6.5198 | -102.3327
6.5235 | -102.3408
6.5272 | -102.3445
6.5308 | -102.3484
6.5345 | -102.3521
6.5382 | -102.3561
6.5418 | -102.3636
6.5455 | -102.3696
6.5492 | -102.3774
6.5528 | -102.3869
6.5565 | -102.3929
6.5601 | -102.396
6.5638 | -102.393
6.5675 | -102.3824
6.5711 | -102.3713
6.5748 | -102.3567
6.5785 | -102.3327
6.5821 | -102.3101
6.5858 | -102.2916
6.5894 | -102.2694
6.5931 | -102.2469
6.5968 | -102.2371
6.6004 | -102.2326
6.6041 | -102.2278
6.6078 | -102.2278
6.6114 | -102.2364
6.6151 | -102.2499
6.6187 | -102.2696
6.6224 | -102.2964
6.6261 | -102.3326
6.6297 | -102.3634
6.6334 | -102.3833
6.6371 | -102.399
6.6407 | -102.4093
6.6444 | -102.4127
6.648 | -102.4074
6.6517 | -102.3957
6.6554 | -102.3816
6.659 | -102.3608
6.6627 | -102.3407
6.6664 | -102.3232
6.67 | -102.3038
6.6737 | -102.2879
6.6773 | -102.2849
6.681 | -102.283
6.6847 | -102.2815
6.6883 | -102.2845
6.692 | -102.2862
6.6957 | -102.2856
6.6993 | -102.285
6.703 | -102.2853
6.7067 | -102.2843
6.7103 | -102.2827
6.714 | -102.2786
6.7176 | -102.2787
6.7213 | -102.2832
6.725 | -102.2932
6.7286 | -102.3129
6.7323 | -102.3382
6.736 | -102.363
6.7396 | -102.3905
6.7433 | -102.4192
6.7469 | -102.4437
6.7506 | -102.4627
6.7543 | -102.4753
6.7579 | -102.4806
6.7616 | -102.4747
6.7653 | -102.462
6.7689 | -102.4456
6.7726 | -102.4291
6.7762 | -102.409
6.7799 | -102.383
6.7836 | -102.3548
6.7872 | -102.3235
6.7909 | -102.2983
6.7946 | -102.2772
6.7982 | -102.2537
6.8019 | -102.2296
6.8055 | -102.2091
6.8092 | -102.1924
6.8129 | -102.1841
6.8165 | -102.1835
6.8202 | -102.183
6.8239 | -102.1797
6.8275 | -102.1744
6.8312 | -102.1657
6.8349 | -102.1582
6.8385 | -102.1481
6.8422 | -102.1276
6.8458 | -102.1051
6.8495 | -102.0796
6.8532 | -102.0494
6.8568 | -102.0139
6.8605 | -101.9856
6.8642 | -101.9568
6.8678 | -101.9309
6.8715 | -101.9013
6.8751 | -101.8732
6.8788 | -101.8496
6.8825 | -101.8313
6.8861 | -101.8179
6.8898 | -101.8014
6.8935 | -101.7843
6.8971 | -101.7748
6.9008 | -101.7714
6.9044 | -101.7775
6.9081 | -101.7881
6.9118 | -101.8024
6.9154 | -101.8244
6.9191 | -101.857
6.9228 | -101.9036
6.9264 | -101.9619
6.9301 | -102.0334
6.9337 | -102.1114
6.9374 | -102.1903
6.9411 | -102.2694
6.9447 | -102.3561
6.9484 | -102.4367
6.9521 | -102.5176
6.9557 | -102.597
6.9594 | -102.669
6.9631 | -102.7211
6.9667 | -102.7678
6.9704 | -102.8092
6.974 | -102.8466
6.9777 | -102.8714
6.9814 | -102.89
6.985 | -102.8973
6.9887 | -102.8991
6.9924 | -102.8934
6.996 | -102.8899
6.9997 | -102.8762
7.0033 | -102.8475
7.007 | -102.8097
7.0107 | -102.766
7.0143 | -102.7183
7.018 | -102.6671
7.0217 | -102.612
7.0253 | -102.5562
7.029 | -102.502
7.0326 | -102.4508
7.0363 | -102.4041
7.04 | -102.367
7.0436 | -102.3368
7.0473 | -102.3161
7.051 | -102.3061
7.0546 | -102.305
7.0583 | -102.3109
7.0619 | -102.3211
7.0656 | -102.3371
7.0693 | -102.3632
7.0729 | -102.3905
7.0766 | -102.4174
7.0803 | -102.4457
7.0839 | -102.4776
7.0876 | -102.5071
7.0912 | -102.5357
7.0949 | -102.5615
7.0986 | -102.582
7.1022 | -102.5958
7.1059 | -102.6124
7.1096 | -102.624
7.1132 | -102.6334
7.1169 | -102.6379
7.1206 | -102.6415
7.1242 | -102.6431
7.1279 | -102.6374
7.1315 | -102.6311
7.1352 | -102.6269
7.1389 | -102.6201
7.1425 | -102.6096
7.1462 | -102.6023
7.1499 | -102.603
7.1535 | -102.6086
7.1572 | -102.615
7.1608 | -102.6305
7.1645 | -102.6498
7.1682 | -102.6735
7.1718 | -102.7008
7.1755 | -102.7223
7.1792 | -102.7439
7.1828 | -102.7673
7.1865 | -102.7856
7.1901 | -102.8017
7.1938 | -102.8146
7.1975 | -102.8191
7.2011 | -102.8172
7.2048 | -102.8156
7.2085 | -102.8076
7.2121 | -102.7902
7.2158 | -102.7697
7.2195 | -102.7464
7.2231 | -102.7193
7.2268 | -102.6819
7.2304 | -102.6397
7.2341 | -102.6018
7.2378 | -102.5603
7.2414 | -102.5161
7.2451 | -102.4829
7.2488 | -102.456
7.2524 | -102.4317
7.2561 | -102.4066
7.2597 | -102.3908
7.2634 | -102.3832
7.2671 | -102.3849
7.2707 | -102.3993
7.2744 | -102.4275
7.2781 | -102.4573
7.2817 | -102.4949
7.2854 | -102.539
7.289 | -102.5812
7.2927 | -102.6228
7.2964 | -102.6615
7.3 | -102.6968
7.3037 | -102.7341
7.3074 | -102.7649
7.311 | -102.7806
7.3147 | -102.7878
7.3183 | -102.7835
7.322 | -102.7765
7.3257 | -102.766
7.3293 | -102.7543
7.333 | -102.7477
7.3367 | -102.74
7.3403 | -102.7299
7.344 | -102.7231
7.3476 | -102.7206
7.3513 | -102.7099
7.355 | -102.6939
7.3586 | -102.6793
7.3623 | -102.6654
7.366 | -102.651
7.3696 | -102.6382
7.3733 | -102.6167
7.377 | -102.5945
7.3806 | -102.5757
7.3843 | -102.5604
7.3879 | -102.5387
7.3916 | -102.5179
7.3953 | -102.5221
7.3989 | -102.5181
7.4026 | -102.5202
7.4063 | -102.5301
7.4099 | -102.5455
7.4136 | -102.5625
7.4172 | -102.5886
7.4209 | -102.6255
7.4246 | -102.6675
7.4282 | -102.7153
7.4319 | -102.7608
7.4356 | -102.8058
7.4392 | -102.851
7.4429 | -102.8963
7.4465 | -102.9461
7.4502 | -102.991
7.4539 | -103.0314
7.4575 | -103.0763
7.4612 | -103.1209
7.4649 | -103.1612
7.4685 | -103.2011
7.4722 | -103.2397
7.4758 | -103.2707
7.4795 | -103.3015
7.4832 | -103.3291
7.4868 | -103.3486
7.4905 | -103.3552
7.4942 | -103.3574
7.4978 | -103.352
7.5015 | -103.3302
7.5052 | -103.2923
7.5088 | -103.2485
7.5125 | -103.2026
7.5161 | -103.1525
7.5198 | -103.0971
7.5235 | -103.0517
7.5271 | -103.0101
7.5308 | -102.9707
7.5345 | -102.9426
7.5381 | -102.9188
7.5418 | -102.8979
7.5454 | -102.879
7.5491 | -102.8584
7.5528 | -102.8353
7.5564 | -102.811
7.5601 | -102.7812
7.5638 | -102.7528
7.5674 | -102.7214
7.5711 | -102.6879
7.5747 | -102.6612
7.5784 | -102.6379
7.5821 | -102.6131
7.5857 | -102.5979
7.5894 | -102.591
7.5931 | -102.5912
7.5967 | -102.6011
7.6004 | -102.6136
7.604 | -102.6345
7.6077 | -102.6592
7.6114 | -102.6835
7.615 | -102.7103
7.6187 | -102.7347
7.6224 | -102.7586
7.626 | -102.7813
7.6297 | -102.7978
7.6334 | -102.8121
7.637 | -102.8187
7.6407 | -102.8253
7.6443 | -102.8302
7.648 | -102.8277
7.6517 | -102.8265
7.6553 | -102.8214
7.659 | -102.8228
7.6627 | -102.8352
7.6663 | -102.8503
7.67 | -102.8678
7.6736 | -102.8928
7.6773 | -102.92
7.681 | -102.9503
7.6846 | -102.9788
7.6883 | -103.001
7.692 | -103.0233
7.6956 | -103.0384
7.6993 | -103.0403
7.7029 | -103.0355
7.7066 | -103.0259
7.7103 | -103.0166
7.7139 | -103.0053
7.7176 | -102.9844
7.7213 | -102.9635
7.7249 | -102.9464
7.7286 | -102.9259
7.7322 | -102.9096
7.7359 | -102.8925
7.7396 | -102.8757
7.7432 | -102.8702
7.7469 | -102.8718
7.7506 | -102.8762
7.7542 | -102.8907
7.7579 | -102.9096
7.7615 | -102.9288
7.7652 | -102.9531
7.7689 | -102.9813
7.7725 | -102.9996
7.7762 | -103.0119
7.7799 | -103.0298
7.7835 | -103.0444
7.7872 | -103.0473
7.7909 | -103.0472
7.7945 | -103.0437
7.7982 | -103.042
7.8018 | -103.042
7.8055 | -103.0438
7.8092 | -103.0505
7.8128 | -103.063
7.8165 | -103.0806
7.8202 | -103.103
7.8238 | -103.131
7.8275 | -103.1687
7.8311 | -103.2162
7.8348 | -103.271
7.8385 | -103.3274
7.8421 | -103.3871
7.8458 | -103.4426
7.8495 | -103.498
7.8531 | -103.5517
7.8568 | -103.5954
7.8604 | -103.6304
7.8641 | -103.6584
7.8678 | -103.6722
7.8714 | -103.6768
7.8751 | -103.6756
7.8788 | -103.6643
7.8824 | -103.643
7.8861 | -103.614
7.8897 | -103.5815
7.8934 | -103.5442
7.8971 | -103.5059
7.9007 | -103.4772
7.9044 | -103.4488
7.9081 | -103.421
7.9117 | -103.3894
7.9154 | -103.3607
7.9191 | -103.3282
7.9227 | -103.3008
7.9264 | -103.2985
7.93 | -103.2856
7.9337 | -103.2765
7.9374 | -103.2746
7.941 | -103.2801
7.9447 | -103.2914
7.9484 | -103.3082
7.952 | -103.3249
7.9557 | -103.3484
7.9593 | -103.3804
7.963 | -103.4145
7.9667 | -103.4494
7.9703 | -103.4852
7.974 | -103.5225
7.9777 | -103.56
7.9813 | -103.5945
7.985 | -103.6278
7.9886 | -103.6576
7.9923 | -103.6868
7.996 | -103.7067
7.9996 | -103.7249
8.0033 | -103.7416
8.007 | -103.7508
8.0106 | -103.7535
8.0143 | -103.7527
8.0179 | -103.7432
8.0216 | -103.7311
8.0253 | -103.721
8.0289 | -103.7131
8.0326 | -103.7078
8.0363 | -103.7033
8.0399 | -103.6904
8.0436 | -103.6725
8.0473 | -103.655
8.0509 | -103.6385
8.0546 | -103.6184
8.0582 | -103.5961
8.0619 | -103.5758
8.0656 | -103.5541
8.0692 | -103.5321
8.0729 | -103.5133
8.0766 | -103.5006
8.0802 | -103.4854
8.0839 | -103.4719
8.0875 | -103.4582
8.0912 | -103.4452
8.0949 | -103.4353
8.0985 | -103.4278
8.1022 | -103.4258
8.1059 | -103.4228
8.1095 | -103.4208
8.1132 | -103.4122
8.1168 | -103.4051
8.1205 | -103.3967
8.1242 | -103.3863
8.1278 | -103.374
8.1315 | -103.3628
8.1352 | -103.3469
8.1388 | -103.3307
8.1425 | -103.3126
8.1461 | -103.2929
8.1498 | -103.2731
8.1535 | -103.2555
8.1571 | -103.2415
8.1608 | -103.2306
8.1645 | -103.2224
8.1681 | -103.2139
8.1718 | -103.2072
8.1754 | -103.2035
8.1791 | -103.2019
8.1828 | -103.1986
8.1864 | -103.2017
8.1901 | -103.208
8.1938 | -103.2152
8.1974 | -103.2252
8.2011 | -103.2421
8.2048 | -103.2643
8.2084 | -103.2865
8.2121 | -103.3088
8.2157 | -103.3311
8.2194 | -103.3547
8.2231 | -103.38
8.2267 | -103.4012
8.2304 | -103.4196
8.2341 | -103.4414
8.2377 | -103.4571
8.2414 | -103.464
8.245 | -103.4673
8.2487 | -103.4706
8.2524 | -103.4695
8.256 | -103.4663
8.2597 | -103.4596
8.2634 | -103.4488
8.267 | -103.4366
8.2707 | -103.4282
8.2743 | -103.4218
8.278 | -103.4251
8.2817 | -103.4309
8.2853 | -103.4335
8.289 | -103.4421
8.2927 | -103.4558
8.2963 | -103.4742
8.3 | -103.4925
8.3036 | -103.5116
8.3073 | -103.5297
8.311 | -103.5534
8.3146 | -103.5792
8.3183 | -103.6011
8.322 | -103.6204
8.3256 | -103.6361
8.3293 | -103.6476
8.333 | -103.6564
8.3366 | -103.659
8.3403 | -103.654
8.3439 | -103.6396
8.3476 | -103.6261
8.3513 | -103.6067
8.3549 | -103.5793
8.3586 | -103.5547
8.3623 | -103.5315
8.3659 | -103.508
8.3696 | -103.4794
8.3732 | -103.455
8.3769 | -103.4427
8.3806 | -103.432
8.3842 | -103.429
8.3879 | -103.428
8.3916 | -103.4294
8.3952 | -103.4367
8.3989 | -103.4499
8.4025 | -103.4739
8.4062 | -103.5034
8.4099 | -103.5338
8.4135 | -103.5652
8.4172 | -103.5972
8.4209 | -103.6297
8.4245 | -103.6606
8.4282 | -103.6919
8.4318 | -103.7212
8.4355 | -103.7468
8.4392 | -103.7686
8.4428 | -103.7829
8.4465 | -103.7912
8.4502 | -103.7929
8.4538 | -103.772
8.4575 | -103.7644
8.4612 | -103.7498
8.4648 | -103.7277
8.4685 | -103.7049
8.4721 | -103.6796
8.4758 | -103.6563
8.4795 | -103.6333
8.4831 | -103.6066
8.4868 | -103.584
8.4905 | -103.5666
8.4941 | -103.5595
8.4978 | -103.5513
8.5014 | -103.5408
8.5051 | -103.5388
8.5088 | -103.5419
8.5124 | -103.5508
8.5161 | -103.5606
8.5198 | -103.5745
8.5234 | -103.59
8.5271 | -103.6052
8.5307 | -103.621
8.5344 | -103.6335
8.5381 | -103.643
8.5417 | -103.6502
8.5454 | -103.6549
8.5491 | -103.6585
8.5527 | -103.6571
8.5564 | -103.6493
8.56 | -103.643
8.5637 | -103.6371
8.5674 | -103.6307
8.571 | -103.6261
8.5747 | -103.6231
8.5784 | -103.6156
8.582 | -103.6075
8.5857 | -103.6031
8.5893 | -103.6047
8.593 | -103.6157
8.5967 | -103.6234
8.6003 | -103.6359
8.604 | -103.6521
8.6077 | -103.6771
8.6113 | -103.7049
8.615 | -103.7325
8.6187 | -103.7621
8.6223 | -103.7883
8.626 | -103.8146
8.6296 | -103.8396
8.6333 | -103.8588
8.637 | -103.874
8.6406 | -103.8852
8.6443 | -103.8899
8.648 | -103.886
8.6516 | -103.88
8.6553 | -103.8729
8.6589 | -103.8645
8.6626 | -103.8556
8.6663 | -103.8421
8.6699 | -103.8315
8.6736 | -103.8255
8.6773 | -103.8236
8.6809 | -103.8256
8.6846 | -103.8295
8.6882 | -103.8361
8.6919 | -103.8445
8.6956 | -103.8537
8.6992 | -103.8642
8.7029 | -103.8721
8.7066 | -103.8801
8.7102 | -103.8897
8.7139 | -103.8954
8.7175 | -103.8942
8.7212 | -103.8864
8.7249 | -103.877
8.7285 | -103.8639
8.7322 | -103.8454
8.7359 | -103.8141
8.7395 | -103.7748
8.7432 | -103.7343
8.7469 | -103.6951
8.7505 | -103.6554
8.7542 | -103.6187
8.7578 | -103.5803
8.7615 | -103.5457
8.7652 | -103.5198
8.7688 | -103.4945
8.7725 | -103.4766
8.7762 | -103.4635
8.7798 | -103.4539
8.7835 | -103.4545
8.7871 | -103.4613
8.7908 | -103.4751
8.7945 | -103.4936
8.7981 | -103.5092
8.8018 | -103.524
8.8055 | -103.5406
8.8091 | -103.5555
8.8128 | -103.5743
8.8164 | -103.5927
8.8201 | -103.6135
8.8238 | -103.6392
8.8274 | -103.6643
8.8311 | -103.6872
8.8348 | -103.7083
8.8384 | -103.7276
8.8421 | -103.7472
8.8457 | -103.7661
8.8494 | -103.7849
8.8531 | -103.8032
8.8567 | -103.8082
8.8604 | -103.8094
8.8641 | -103.8155
8.8677 | -103.8147
8.8714 | -103.8057
8.8751 | -103.7853
8.8787 | -103.7609
8.8824 | -103.7345
8.886 | -103.7062
8.8897 | -103.6753
8.8934 | -103.6438
8.897 | -103.623
8.9007 | -103.6015
8.9044 | -103.581
8.908 | -103.5648
8.9117 | -103.5493
8.9153 | -103.538
8.919 | -103.5321
8.9227 | -103.53
8.9263 | -103.5327
8.93 | -103.538
8.9337 | -103.5446
8.9373 | -103.5566
8.941 | -103.569
8.9446 | -103.5835
8.9483 | -103.5976
8.952 | -103.6125
8.9556 | -103.6332
8.9593 | -103.6665
8.963 | -103.7072
8.9666 | -103.7489
8.9703 | -103.7899
8.9739 | -103.836
8.9776 | -103.8858
8.9813 | -103.9483
8.9849 | -104.0035
8.9886 | -104.0654
8.9923 | -104.1223
8.9959 | -104.1758
8.9996 | -104.2348
9.0033 | -104.2913
9.0069 | -104.3419
9.0106 | -104.3889
9.0142 | -104.4303
9.0179 | -104.4716
9.0216 | -104.5059
9.0252 | -104.5266
9.0289 | -104.5411
9.0326 | -104.5549
9.0362 | -104.5623
9.0399 | -104.5581
9.0435 | -104.5436
9.0472 | -104.5224
9.0509 | -104.4899
9.0545 | -104.4517
9.0582 | -104.407
9.0619 | -104.3596
9.0655 | -104.3083
9.0692 | -104.2533
9.0728 | -104.1992
9.0765 | -104.1448
9.0802 | -104.0917
9.0838 | -104.0387
9.0875 | -103.9879
9.0912 | -103.9362
9.0948 | -103.8883
9.0985 | -103.8479
9.1021 | -103.8114
9.1058 | -103.7774
9.1095 | -103.7504
9.1131 | -103.7305
9.1168 | -103.7139
9.1205 | -103.7032
9.1241 | -103.701
9.1278 | -103.7069
9.1314 | -103.7156
9.1351 | -103.7257
9.1388 | -103.7411
9.1424 | -103.7661
9.1461 | -103.7949
9.1498 | -103.8229
9.1534 | -103.8519
9.1571 | -103.8841
9.1608 | -103.9177
9.1644 | -103.9523
9.1681 | -103.9878
9.1717 | -104.0255
9.1754 | -104.059
9.1791 | -104.0938
9.1827 | -104.1272
9.1864 | -104.162
9.1901 | -104.1953
9.1937 | -104.2272
9.1974 | -104.2528
9.201 | -104.2753
9.2047 | -104.2916
9.2084 | -104.3044
9.212 | -104.3132
9.2157 | -104.3166
9.2194 | -104.314
9.223 | -104.3077
9.2267 | -104.2996
9.2303 | -104.2893
9.234 | -104.2784
9.2377 | -104.2654
9.2413 | -104.2464
9.245 | -104.2228
9.2487 | -104.1991
9.2523 | -104.1817
9.256 | -104.1642
9.2596 | -104.1428
9.2633 | -104.122
9.267 | -104.1067
9.2706 | -104.095
9.2743 | -104.0806
9.278 | -104.0645
9.2816 | -104.0498
9.2853 | -104.0433
9.289 | -104.0378
9.2926 | -104.038
9.2963 | -104.0451
9.2999 | -104.0512
9.3036 | -104.0575
9.3073 | -104.0705
9.3109 | -104.0944
9.3146 | -104.1233
9.3183 | -104.1541
9.3219 | -104.1856
9.3256 | -104.2207
9.3292 | -104.2574
9.3329 | -104.297
9.3366 | -104.3364
9.3402 | -104.3773
9.3439 | -104.4149
9.3476 | -104.4483
9.3512 | -104.479
9.3549 | -104.5034
9.3585 | -104.5214
9.3622 | -104.5352
9.3659 | -104.542
9.3695 | -104.5436
9.3732 | -104.5466
9.3769 | -104.5458
9.3805 | -104.5415
9.3842 | -104.5313
9.3878 | -104.5175
9.3915 | -104.5045
9.3952 | -104.4923
9.3988 | -104.4807
9.4025 | -104.4643
9.4062 | -104.446
9.4098 | -104.4307
9.4135 | -104.4208
9.4172 | -104.4142
9.4208 | -104.4069
9.4245 | -104.4032
9.4281 | -104.3993
9.4318 | -104.3949
9.4355 | -104.3931
9.4391 | -104.3958
9.4428 | -104.4013
9.4465 | -104.4057
9.4501 | -104.4115
9.4538 | -104.4227
9.4574 | -104.4328
9.4611 | -104.4446
9.4648 | -104.4585
9.4684 | -104.4716
9.4721 | -104.4849
9.4758 | -104.4974
9.4794 | -104.5058
9.4831 | -104.517
9.4867 | -104.5296
9.4904 | -104.5431
9.4941 | -104.5518
9.4977 | -104.5602
9.5014 | -104.5703
9.5051 | -104.579
9.5087 | -104.5882
9.5124 | -104.5989
9.516 | -104.6126
9.5197 | -104.6249
9.5234 | -104.6387
9.527 | -104.6503
9.5307 | -104.6664
9.5344 | -104.6865
9.538 | -104.7058
9.5417 | -104.7245
9.5453 | -104.7458
9.549 | -104.7643
9.5527 | -104.7782
9.5563 | -104.7867
9.56 | -104.7927
9.5637 | -104.8017
9.5673 | -104.8067
9.571 | -104.815
9.5747 | -104.8224
9.5783 | -104.8251
9.582 | -104.8226
9.5856 | -104.8203
9.5893 | -104.8153
9.593 | -104.8136
9.5966 | -104.8183
9.6003 | -104.8203
9.604 | -104.8234
9.6076 | -104.8282
9.6113 | -104.8337
9.6149 | -104.8398
9.6186 | -104.8482
9.6223 | -104.8557
9.6259 | -104.8615
9.6296 | -104.8662
9.6333 | -104.8749
9.6369 | -104.8803
9.6406 | -104.8796
9.6442 | -104.8757
9.6479 | -104.8691
9.6516 | -104.8646
9.6552 | -104.8545
9.6589 | -104.8413
9.6626 | -104.829
9.6662 | -104.8145
9.6699 | -104.7976
9.6736 | -104.7725
9.6772 | -104.7429
9.6809 | -104.7126
9.6845 | -104.6826
9.6882 | -104.6465
9.6919 | -104.6123
9.6955 | -104.5792
9.6992 | -104.5465
9.7029 | -104.5124
9.7065 | -104.4794
9.7102 | -104.4513
9.7138 | -104.4247
9.7175 | -104.4047
9.7212 | -104.3938
9.7248 | -104.3877
9.7285 | -104.3865
9.7322 | -104.3915
9.7358 | -104.3997
9.7395 | -104.4153
9.7431 | -104.4406
9.7468 | -104.471
9.7505 | -104.5028
9.7541 | -104.532
9.7578 | -104.5657
9.7615 | -104.6006
9.7651 | -104.6316
9.7688 | -104.658
9.7724 | -104.6841
9.7761 | -104.7047
9.7798 | -104.721
9.7834 | -104.7403
9.7871 | -104.7546
9.7908 | -104.7631
9.7944 | -104.7668
9.7981 | -104.7663
9.8017 | -104.7672
9.8054 | -104.7663
9.8091 | -104.761
9.8127 | -104.7582
9.8164 | -104.7588
9.8201 | -104.7581
9.8237 | -104.7522
9.8274 | -104.747
9.8311 | -104.7407
9.8347 | -104.729
9.8384 | -104.718
9.842 | -104.7056
9.8457 | -104.6938
9.8494 | -104.6795
9.853 | -104.6669
9.8567 | -104.6513
9.8604 | -104.6307
9.864 | -104.6108
9.8677 | -104.5924
9.8713 | -104.5741
9.875 | -104.5598
9.8787 | -104.5466
9.8823 | -104.5351
9.886 | -104.5246
9.8897 | -104.5162
9.8933 | -104.5151
9.897 | -104.5208
9.9006 | -104.5267
9.9043 | -104.5328
9.908 | -104.5396
9.9116 | -104.5485
9.9153 | -104.5552
9.919 | -104.5593
9.9226 | -104.565
9.9263 | -104.5702
9.9299 | -104.5722
9.9336 | -104.5747
9.9373 | -104.5735
9.9409 | -104.5724
9.9446 | -104.5647
9.9483 | -104.5584
9.9519 | -104.5561
9.9556 | -104.5473
9.9593 | -104.5309
9.9629 | -104.5149
9.9666 | -104.5022
9.9702 | -104.488
9.9739 | -104.4739
9.9776 | -104.4609
9.9812 | -104.4486
9.9849 | -104.4368
9.9886 | -104.4249
9.9922 | -104.4162
9.9959 | -104.413
9.9995 | -104.4171
10.0032 | -104.4249
10.0069 | -104.431
10.0105 | -104.4424
10.0142 | -104.4555
10.0179 | -104.4643
10.0215 | -104.4727
10.0252 | -104.477
10.0288 | -104.4821
10.0325 | -104.4782
10.0362 | -104.4703
10.0398 | -104.4562
10.0435 | -104.452
10.0472 | -104.4445
10.0508 | -104.4362
10.0545 | -104.4276
10.0581 | -104.4181
10.0618 | -104.412
10.0655 | -104.411
10.0691 | -104.4082
10.0728 | -104.3992
10.0765 | -104.398
10.0801 | -104.3966
10.0838 | -104.3956
10.0875 | -104.3955
10.0911 | -104.3968
10.0948 | -104.3951
10.0984 | -104.3936
10.1021 | -104.3963
10.1058 | -104.4055
10.1094 | -104.4154
10.1131 | -104.4206
10.1168 | -104.4287
10.1204 | -104.4417
10.1241 | -104.4568
10.1277 | -104.4705
10.1314 | -104.485
10.1351 | -104.5024
10.1387 | -104.5209
10.1424 | -104.5345
10.1461 | -104.5468
10.1497 | -104.5591
10.1534 | -104.5709
10.157 | -104.579
10.1607 | -104.5876
10.1644 | -104.5992
10.168 | -104.6098
10.1717 | -104.6216
10.1754 | -104.6344
10.179 | -104.6507
10.1827 | -104.6646
10.1863 | -104.6743
10.19 | -104.6809
10.1937 | -104.6881
10.1973 | -104.6976
10.201 | -104.7115
10.2047 | -104.7298
10.2083 | -104.7472
10.212 | -104.7637
10.2156 | -104.7781
10.2193 | -104.7945
10.223 | -104.8116
10.2266 | -104.8243
10.2303 | -104.8395
10.234 | -104.8498
10.2376 | -104.8601
10.2413 | -104.8718
10.245 | -104.879
10.2486 | -104.8852
10.2523 | -104.8879
10.2559 | -104.8816
10.2596 | -104.8738
10.2633 | -104.8706
10.2669 | -104.862
10.2706 | -104.8502
10.2743 | -104.8406
10.2779 | -104.8295
10.2816 | -104.8214
10.2852 | -104.8114
10.2889 | -104.8
10.2926 | -104.795
10.2962 | -104.7931
10.2999 | -104.7948
10.3036 | -104.7955
10.3072 | -104.7951
10.3109 | -104.7993
10.3145 | -104.8032
10.3182 | -104.8089
10.3219 | -104.8169
10.3255 | -104.8251
10.3292 | -104.8356
10.3329 | -104.8487
10.3365 | -104.8584
10.3402 | -104.8653
10.3438 | -104.8656
10.3475 | -104.8656
10.3512 | -104.8665
10.3548 | -104.864
10.3585 | -104.8625
10.3622 | -104.8601
10.3658 | -104.8594
10.3695 | -104.8625
10.3732 | -104.869
10.3768 | -104.874
10.3805 | -104.8748
10.3841 | -104.8769
10.3878 | -104.8782
10.3915 | -104.8829
10.3951 | -104.8888
10.3988 | -104.8948
10.4025 | -104.9035
10.4061 | -104.9145
10.4098 | -104.9267
10.4134 | -104.9327
10.4171 | -104.9388
10.4208 | -104.9445
10.4244 | -104.9498
10.4281 | -104.9569
10.4318 | -104.9638
10.4354 | -104.9678
10.4391 | -104.9718
10.4427 | -104.9736
10.4464 | -104.9751
10.4501 | -104.9775
10.4537 | -104.9836
10.4574 | -104.9917
10.4611 | -104.993
10.4647 | -104.9974
10.4684 | -105.0042
10.472 | -105.0072
10.4757 | -105.008
10.4794 | -105.0071
10.483 | -105.0064
10.4867 | -105.0091
10.4904 | -105.0106
10.494 | -105.0145
10.4977 | -105.0143
10.5014 | -105.0104
10.505 | -105.0065
10.5087 | -105.0003
10.5123 | -104.9966
10.516 | -104.9942
10.5197 | -104.9901
10.5233 | -104.9864
10.527 | -104.9826
10.5307 | -104.9812
10.5343 | -104.9852
10.538 | -104.9833
10.5416 | -104.974
10.5453 | -104.9706
10.549 | -104.9664
10.5526 | -104.9608
10.5563 | -104.9563
10.56 | -104.9518
10.5636 | -104.9481
10.5673 | -104.9509
10.5709 | -104.955
10.5746 | -104.9597
10.5783 | -104.9629
10.5819 | -104.9664
10.5856 | -104.9697
10.5893 | -104.9779
10.5929 | -104.9887
10.5966 | -104.9991
10.6002 | -105.0077
10.6039 | -105.0139
10.6076 | -105.0231
10.6112 | -105.0331
10.6149 | -105.0454
10.6186 | -105.0561
10.6222 | -105.0659
10.6259 | -105.0807
10.6296 | -105.0958
10.6332 | -105.1104
10.6369 | -105.1177
10.6405 | -105.1246
10.6442 | -105.1341
10.6479 | -105.1424
10.6515 | -105.1515
10.6552 | -105.1562
10.6589 | -105.1627
10.6625 | -105.1713
10.6662 | -105.1753
10.6698 | -105.1769
10.6735 | -105.1783
10.6772 | -105.1782
10.6808 | -105.1821
10.6845 | -105.1848
10.6882 | -105.1839
10.6918 | -105.1877
10.6955 | -105.1918
10.6991 | -105.189
10.7028 | -105.1835
10.7065 | -105.1812
10.7101 | -105.1772
10.7138 | -105.1714
10.7175 | -105.1694
10.7211 | -105.165
10.7248 | -105.1569
10.7284 | -105.148
10.7321 | -105.1427
10.7358 | -105.1373
10.7394 | -105.134
10.7431 | -105.1336
10.7468 | -105.128
10.7504 | -105.1185
10.7541 | -105.1148
10.7577 | -105.1122
10.7614 | -105.1152
10.7651 | -105.1131
10.7687 | -105.1127
10.7724 | -105.1119
10.7761 | -105.1091
10.7797 | -105.1095
10.7834 | -105.1094
10.7871 | -105.1054
10.7907 | -105.0994
10.7944 | -105.0963
10.798 | -105.0918
10.8017 | -105.083
10.8054 | -105.0706
10.809 | -105.0561
10.8127 | -105.0425
10.8164 | -105.0245
10.82 | -105.0039
10.8237 | -104.985
10.8273 | -104.9678
10.831 | -104.9512
10.8347 | -104.938
10.8383 | -104.9209
10.842 | -104.9053
10.8457 | -104.8945
10.8493 | -104.8866
10.853 | -104.8831
10.8566 | -104.8818
10.8603 | -104.8825
10.864 | -104.8838
10.8676 | -104.8909
10.8713 | -104.9017
10.875 | -104.9111
10.8786 | -104.9222
10.8823 | -104.9368
10.8859 | -104.9517
10.8896 | -104.9672
10.8933 | -104.9838
10.8969 | -104.9945
10.9006 | -105.0002
10.9043 | -105.0113
10.9079 | -105.0223
10.9116 | -105.0289
10.9153 | -105.0344
10.9189 | -105.0352
10.9226 | -105.0374
10.9262 | -105.0379
10.9299 | -105.037
10.9336 | -105.0366
10.9372 | -105.0356
10.9409 | -105.036
10.9446 | -105.0417
10.9482 | -105.0448
10.9519 | -105.0458
10.9555 | -105.0474
10.9592 | -105.0542
10.9629 | -105.0658
10.9665 | -105.0784
10.9702 | -105.0918
10.9739 | -105.1107
10.9775 | -105.1315
10.9812 | -105.1512
10.9848 | -105.1732
10.9885 | -105.1977
10.9922 | -105.2233
10.9958 | -105.2524
10.9995 | -105.2832
11.0032 | -105.3161
11.0068 | -105.3445
11.0105 | -105.3719
11.0141 | -105.3986
11.0178 | -105.4237
11.0215 | -105.443
11.0251 | -105.456
11.0288 | -105.4655
11.0325 | -105.4702
11.0361 | -105.4699
11.0398 | -105.4653
11.0434 | -105.4565
11.0471 | -105.4391
11.0508 | -105.4146
11.0544 | -105.3933
11.0581 | -105.3709
11.0618 | -105.34
11.0654 | -105.3088
11.0691 | -105.2768
11.0728 | -105.2473
11.0764 | -105.2191
11.0801 | -105.1943
11.0837 | -105.172
11.0874 | -105.1548
11.0911 | -105.1417
11.0947 | -105.1582
11.0984 | -105.1597
11.1021 | -105.1622
11.1057 | -105.1676
11.1094 | -105.1826
11.113 | -105.2006
11.1167 | -105.2198
11.1204 | -105.2406
11.124 | -105.2645
11.1277 | -105.2899
11.1314 | -105.3148
11.135 | -105.338
11.1387 | -105.3614
11.1423 | -105.3848
11.146 | -105.4032
11.1497 | -105.4194
11.1533 | -105.4328
11.157 | -105.4408
11.1607 | -105.4441
11.1643 | -105.4435
11.168 | -105.4448
11.1716 | -105.4427
11.1753 | -105.4392
11.179 | -105.4294
11.1826 | -105.4156
11.1863 | -105.3972
11.19 | -105.3721
11.1936 | -105.3439
11.1973 | -105.3176
11.201 | -105.2889
11.2046 | -105.2585
11.2083 | -105.2312
11.2119 | -105.2061
11.2156 | -105.178
11.2193 | -105.1497
11.2229 | -105.1248
11.2266 | -105.1023
11.2303 | -105.0805
11.2339 | -105.0574
11.2376 | -105.0421
11.2412 | -105.0298
11.2449 | -105.02
11.2486 | -105.0127
11.2522 | -105.0083
11.2559 | -105.0033
11.2596 | -105.0035
11.2632 | -105.0066
11.2669 | -105.0145
11.2705 | -105.0248
11.2742 | -105.0349
11.2779 | -105.0473
11.2815 | -105.0598
11.2852 | -105.0706
11.2889 | -105.0867
11.2925 | -105.1035
11.2962 | -105.1212
11.2998 | -105.1387
11.3035 | -105.1529
11.3072 | -105.1646
11.3108 | -105.1769
11.3145 | -105.1872
11.3182 | -105.1983
11.3218 | -105.2111
11.3255 | -105.2203
11.3292 | -105.226
11.3328 | -105.231
11.3365 | -105.2372
11.3401 | -105.2447
11.3438 | -105.2538
11.3475 | -105.2634
11.3511 | -105.2739
11.3548 | -105.2862
11.3585 | -105.2999
11.3621 | -105.3074
11.3658 | -105.3134
11.3694 | -105.3239
11.3731 | -105.3315
11.3768 | -105.3378
11.3804 | -105.3453
11.3841 | -105.3532
11.3878 | -105.3601
11.3914 | -105.3649
11.3951 | -105.3739
11.3987 | -105.3813
11.4024 | -105.3831
11.4061 | -105.3882
11.4097 | -105.3923
11.4134 | -105.3954
11.4171 | -105.3969
11.4207 | -105.3957
11.4244 | -105.3957
11.428 | -105.3944
11.4317 | -105.3904
11.4354 | -105.3856
11.439 | -105.3792
11.4427 | -105.3735
11.4464 | -105.3671
11.45 | -105.359
11.4537 | -105.3478
11.4574 | -105.3335
11.461 | -105.3199
11.4647 | -105.3075
11.4683 | -105.2993
11.472 | -105.2901
11.4757 | -105.28
11.4793 | -105.2651
11.483 | -105.2489
11.4867 | -105.2343
11.4903 | -105.2218
11.494 | -105.2097
11.4976 | -105.1983
11.5013 | -105.1863
11.505 | -105.1708
11.5086 | -105.1584
11.5123 | -105.1493
11.516 | -105.1383
11.5196 | -105.1262
11.5233 | -105.1154
11.5269 | -105.1054
11.5306 | -105.1003
11.5343 | -105.0945
11.5379 | -105.0923
11.5416 | -105.0931
11.5453 | -105.0933
11.5489 | -105.0939
11.5526 | -105.0946
11.5562 | -105.098
11.5599 | -105.1021
11.5636 | -105.1067
11.5672 | -105.1112
11.5709 | -105.1095
11.5746 | -105.1066
11.5782 | -105.1085
11.5819 | -105.1119
11.5855 | -105.1146
11.5892 | -105.1142
11.5929 | -105.1154
11.5965 | -105.119
11.6002 | -105.1233
11.6039 | -105.1289
11.6075 | -105.133
11.6112 | -105.1416
11.6149 | -105.1527
11.6185 | -105.1629
11.6222 | -105.1733
11.6258 | -105.188
11.6295 | -105.2017
11.6332 | -105.2123
11.6368 | -105.2216
11.6405 | -105.2322
11.6442 | -105.2449
11.6478 | -105.2617
11.6515 | -105.2735
11.6551 | -105.282
11.6588 | -105.2904
11.6625 | -105.2965
11.6661 | -105.3013
11.6698 | -105.304
11.6735 | -105.309
11.6771 | -105.311
11.6808 | -105.3102
11.6844 | -105.311
11.6881 | -105.3122
11.6918 | -105.3131
11.6954 | -105.316
11.6991 | -105.3171
11.7028 | -105.3179
11.7064 | -105.3202
11.7101 | -105.323
11.7137 | -105.3287
11.7174 | -105.3358
11.7211 | -105.3458
11.7247 | -105.3553
11.7284 | -105.3677
11.7321 | -105.3846
11.7357 | -105.3986
11.7394 | -105.4171
11.7431 | -105.4375
11.7467 | -105.4516
11.7504 | -105.4653
11.754 | -105.4808
11.7577 | -105.4948
11.7614 | -105.5083
11.765 | -105.5252
11.7687 | -105.5384
11.7724 | -105.5485
11.776 | -105.5569
11.7797 | -105.5632
11.7833 | -105.5689
11.787 | -105.573
11.7907 | -105.5788
11.7943 | -105.5862
11.798 | -105.5837
11.8017 | -105.5778
11.8053 | -105.5739
11.809 | -105.5672
11.8126 | -105.5611
11.8163 | -105.5545
11.82 | -105.5455
11.8236 | -105.5374
11.8273 | -105.5302
11.831 | -105.5236
11.8346 | -105.5159
11.8383 | -105.5071
11.8419 | -105.4989
11.8456 | -105.4884
11.8493 | -105.478
11.8529 | -105.4674
11.8566 | -105.4579
11.8603 | -105.4507
11.8639 | -105.4413
11.8676 | -105.4313
11.8713 | -105.4261
11.8749 | -105.42
11.8786 | -105.4095
11.8822 | -105.3981
11.8859 | -105.3901
11.8896 | -105.3823
11.8932 | -105.376
11.8969 | -105.3696
11.9006 | -105.3592
11.9042 | -105.3505
11.9079 | -105.3454
11.9115 | -105.3391
11.9152 | -105.3349
11.9189 | -105.3314
11.9225 | -105.3282
11.9262 | -105.3266
11.9299 | -105.3303
11.9335 | -105.3359
11.9372 | -105.3384
11.9408 | -105.3446
11.9445 | -105.3557
11.9482 | -105.3669
11.9518 | -105.3786
11.9555 | -105.3925
11.9592 | -105.41
11.9628 | -105.43
11.9665 | -105.4493
11.9701 | -105.472
11.9738 | -105.4951
11.9775 | -105.518
11.9811 | -105.5421
11.9848 | -105.5618
11.9885 | -105.5776
11.9921 | -105.5935
11.9958 | -105.612
11.9995 | -105.63
12.0031 | -105.6401
12.0068 | -105.6518
12.0104 | -105.6633
12.0141 | -105.6757
12.0178 | -105.686
12.0214 | -105.6944
12.0251 | -105.6985
12.0288 | -105.7
12.0324 | -105.7007
12.0361 | -105.7021
12.0397 | -105.703
12.0434 | -105.6973
12.0471 | -105.6903
12.0507 | -105.6885
12.0544 | -105.6843
12.0581 | -105.6749
12.0617 | -105.6672
12.0654 | -105.6582
12.069 | -105.6475
12.0727 | -105.6383
12.0764 | -105.628
12.08 | -105.6171
12.0837 | -105.6061
12.0874 | -105.5921
12.091 | -105.5762
12.0947 | -105.561
12.0983 | -105.5452
12.102 | -105.5305
12.1057 | -105.5154
12.1093 | -105.5058
12.113 | -105.4975
12.1167 | -105.4881
12.1203 | -105.4801
12.124 | -105.4726
12.1276 | -105.4634
12.1313 | -105.4545
12.135 | -105.4482
12.1386 | -105.4485
12.1423 | -105.4492
12.146 | -105.4526
12.1496 | -105.4532
12.1533 | -105.4666
12.157 | -105.4724
12.1606 | -105.4809
12.1643 | -105.4927
12.1679 | -105.5084
12.1716 | -105.521
12.1753 | -105.5283
12.1789 | -105.5423
12.1826 | -105.5549
12.1863 | -105.5661
12.1899 | -105.574
12.1936 | -105.5826
12.1972 | -105.5946
12.2009 | -105.606
12.2046 | -105.6147
12.2082 | -105.6203
12.2119 | -105.6231
12.2156 | -105.6263
12.2192 | -105.6287
12.2229 | -105.6283
12.2265 | -105.6213
12.2302 | -105.6152
12.2339 | -105.6048
12.2375 | -105.5914
12.2412 | -105.5748
12.2449 | -105.5576
12.2485 | -105.5426
12.2522 | -105.5259
12.2558 | -105.5089
12.2595 | -105.4894
12.2632 | -105.4696
12.2668 | -105.4524
12.2705 | -105.4343
12.2742 | -105.4158
12.2778 | -105.3982
12.2815 | -105.3817
12.2852 | -105.3654
12.2888 | -105.3505
12.2925 | -105.337
12.2961 | -105.3257
12.2998 | -105.3169
12.3035 | -105.3139
12.3071 | -105.3113
12.3108 | -105.308
12.3145 | -105.3041
12.3181 | -105.3059
12.3218 | -105.3117
12.3254 | -105.3202
12.3291 | -105.3302
12.3328 | -105.3421
12.3364 | -105.3567
12.3401 | -105.3723
12.3438 | -105.3886
12.3474 | -105.4018
12.3511 | -105.4157
12.3547 | -105.4342
12.3584 | -105.4516
12.3621 | -105.4671
12.3657 | -105.481
12.3694 | -105.4928
12.3731 | -105.5008
12.3767 | -105.5017
12.3804 | -105.5022
12.384 | -105.5003
12.3877 | -105.4952
12.3914 | -105.487
12.395 | -105.473
12.3987 | -105.4602
12.4024 | -105.4503
12.406 | -105.4398
12.4097 | -105.425
12.4134 | -105.4086
12.417 | -105.3904
12.4207 | -105.3723
12.4243 | -105.3537
12.428 | -105.3338
12.4317 | -105.3188
12.4353 | -105.3058
12.439 | -105.2921
12.4427 | -105.2811
12.4463 | -105.2731
12.45 | -105.2663
12.4536 | -105.2632
12.4573 | -105.2624
12.461 | -105.2644
12.4646 | -105.2653
12.4683 | -105.2667
12.472 | -105.2699
12.4756 | -105.2752
12.4793 | -105.2782
12.4829 | -105.2799
12.4866 | -105.2816
12.4903 | -105.2864
12.4939 | -105.2929
12.4976 | -105.296
12.5013 | -105.2964
12.5049 | -105.2973
12.5086 | -105.2955
12.5122 | -105.2969
12.5159 | -105.2995
12.5196 | -105.3032
12.5232 | -105.3055
12.5269 | -105.3088
12.5306 | -105.3149
12.5342 | -105.3208
12.5379 | -105.3251
12.5416 | -105.3296
12.5452 | -105.3364
12.5489 | -105.344
12.5525 | -105.3501
12.5562 | -105.3596
12.5599 | -105.3715
12.5635 | -105.3855
12.5672 | -105.4006
12.5709 | -105.4147
12.5745 | -105.4307
12.5782 | -105.4465
12.5818 | -105.46
12.5855 | -105.4731
12.5892 | -105.4873
12.5928 | -105.5053
12.5965 | -105.525
12.6002 | -105.5464
12.6038 | -105.5673
12.6075 | -105.5858
12.6111 | -105.6028
12.6148 | -105.6194
12.6185 | -105.6402
12.6221 | -105.6568
12.6258 | -105.6714
12.6295 | -105.6851
12.6331 | -105.6969
12.6368 | -105.7062
12.6404 | -105.7164
12.6441 | -105.7306
12.6478 | -105.7444
12.6514 | -105.752
12.6551 | -105.757
12.6588 | -105.7663
12.6624 | -105.7718
12.6661 | -105.7724
12.6698 | -105.7713
12.6734 | -105.77
12.6771 | -105.7723
12.6807 | -105.7669
12.6844 | -105.7689
12.6881 | -105.7641
12.6917 | -105.759
12.6954 | -105.7513
12.6991 | -105.7444
12.7027 | -105.7381
12.7064 | -105.7319
12.71 | -105.7279
12.7137 | -105.7226
12.7174 | -105.7168
12.721 | -105.7132
12.7247 | -105.7101
12.7284 | -105.7059
12.732 | -105.707
12.7357 | -105.7124
12.7393 | -105.7152
12.743 | -105.7174
12.7467 | -105.7212
12.7503 | -105.7276
12.754 | -105.7396
12.7577 | -105.7496
12.7613 | -105.7604
12.765 | -105.7729
12.7686 | -105.7904
12.7723 | -105.8097
12.776 | -105.8259
12.7796 | -105.8396
12.7833 | -105.8537
12.787 | -105.87
12.7906 | -105.8846
12.7943 | -105.8972
12.7979 | -105.909
12.8016 | -105.9164
12.8053 | -105.9239
12.8089 | -105.9294
12.8126 | -105.9322
12.8163 | -105.9317
12.8199 | -105.9281
12.8236 | -105.9178
12.8273 | -105.9033
12.8309 | -105.8822
12.8346 | -105.8668
12.8382 | -105.8498
12.8419 | -105.8251
12.8456 | -105.799
12.8492 | -105.7766
12.8529 | -105.7493
12.8566 | -105.7208
12.8602 | -105.69
12.8639 | -105.6602
12.8675 | -105.6333
12.8712 | -105.6028
12.8749 | -105.5762
12.8785 | -105.5512
12.8822 | -105.5235
12.8859 | -105.4963
12.8895 | -105.4709
12.8932 | -105.4463
12.8968 | -105.4267
12.9005 | -105.4134
12.9042 | -105.4023
12.9078 | -105.3909
12.9115 | -105.3831
12.9152 | -105.3788
12.9188 | -105.3769
12.9225 | -105.3774
12.9261 | -105.3814
12.9298 | -105.3874
12.9335 | -105.3943
12.9371 | -105.4041
12.9408 | -105.4177
12.9445 | -105.4361
12.9481 | -105.4524
12.9518 | -105.4691
12.9555 | -105.4804
12.9591 | -105.492
12.9628 | -105.5052
12.9664 | -105.5194
12.9701 | -105.5299
12.9738 | -105.5374
12.9774 | -105.5432
12.9811 | -105.5477
12.9848 | -105.5499
12.9884 | -105.5496
12.9921 | -105.5511
12.9957 | -105.5477
12.9994 | -105.5426
13.0031 | -105.5386
13.0067 | -105.5344
13.0104 | -105.5322
13.0141 | -105.5274
13.0177 | -105.5259
13.0214 | -105.5188
13.025 | -105.511
13.0287 | -105.5057
13.0324 | -105.5016
13.036 | -105.4957
13.0397 | -105.4907
13.0434 | -105.4875
13.047 | -105.4876
13.0507 | -105.4889
13.0543 | -105.4932
13.058 | -105.5006
13.0617 | -105.5093
13.0653 | -105.5159
13.069 | -105.5252
13.0727 | -105.5388
13.0763 | -105.5512
13.08 | -105.5626
13.0837 | -105.5694
13.0873 | -105.5745
13.091 | -105.5803
13.0946 | -105.5878
13.0983 | -105.5947
13.102 | -105.6015
13.1056 | -105.6067
13.1093 | -105.6089
13.113 | -105.612
13.1166 | -105.6117
13.1203 | -105.6121
13.1239 | -105.6094
13.1276 | -105.6041
13.1313 | -105.5985
13.1349 | -105.5936
13.1386 | -105.588
13.1423 | -105.5892
13.1459 | -105.5902
13.1496 | -105.5924
13.1532 | -105.5942
13.1569 | -105.6003
13.1606 | -105.6089
13.1642 | -105.6203
13.1679 | -105.6317
13.1716 | -105.6417
13.1752 | -105.6574
13.1789 | -105.6714
13.1825 | -105.6846
13.1862 | -105.7
13.1899 | -105.7154
13.1935 | -105.7327
13.1972 | -105.7503
13.2009 | -105.7645
13.2045 | -105.7772
13.2082 | -105.7855
13.2118 | -105.7946
13.2155 | -105.8023
13.2192 | -105.8102
13.2228 | -105.8178
13.2265 | -105.8244
13.2302 | -105.829
13.2338 | -105.829
13.2375 | -105.829
13.2412 | -105.8256
13.2448 | -105.8183
13.2485 | -105.8111
13.2521 | -105.8001
13.2558 | -105.7896
13.2595 | -105.7776
13.2631 | -105.7575
13.2668 | -105.7421
13.2705 | -105.7241
13.2741 | -105.704
13.2778 | -105.6832
13.2814 | -105.6601
13.2851 | -105.6344
13.2888 | -105.6055
13.2924 | -105.5751
13.2961 | -105.5448
13.2998 | -105.5173
13.3034 | -105.4844
13.3071 | -105.4529
13.3107 | -105.4225
13.3144 | -105.3943
13.3181 | -105.3668
13.3217 | -105.3378
13.3254 | -105.311
13.3291 | -105.2868
13.3327 | -105.2571
13.3364 | -105.2297
13.34 | -105.2034
13.3437 | -105.1758
13.3474 | -105.1514
13.351 | -105.1305
13.3547 | -105.1076
13.3584 | -105.0839
13.362 | -105.0635
13.3657 | -105.0458
13.3694 | -105.031
13.373 | -105.0212
13.3767 | -105.0133
13.3803 | -105.0058
13.384 | -104.9984
13.3877 | -104.9935
13.3913 | -104.9939
13.395 | -104.996
13.3987 | -104.9978
13.4023 | -105.0033
13.406 | -105.0114
13.4096 | -105.0213
13.4133 | -105.0308
13.417 | -105.0399
13.4206 | -105.0501
13.4243 | -105.0615
13.428 | -105.071
13.4316 | -105.0795
13.4353 | -105.0932
13.4389 | -105.1066
13.4426 | -105.1193
13.4463 | -105.1345
13.4499 | -105.1497
13.4536 | -105.1624
13.4573 | -105.1771
13.4609 | -105.1871
13.4646 | -105.2047
13.4682 | -105.2208
13.4719 | -105.2333
13.4756 | -105.2478
13.4792 | -105.2646
13.4829 | -105.2847
13.4866 | -105.3044
13.4902 | -105.3219
13.4939 | -105.3379
13.4976 | -105.3603
13.5012 | -105.3824
13.5049 | -105.3991
13.5085 | -105.4162
13.5122 | -105.4331
13.5159 | -105.4489
13.5195 | -105.4645
13.5232 | -105.4778
13.5269 | -105.4904
13.5305 | -105.5008
13.5342 | -105.5107
13.5378 | -105.5204
13.5415 | -105.5263
13.5452 | -105.5261
13.5488 | -105.5254
13.5525 | -105.5271
13.5562 | -105.5271
13.5598 | -105.5202
13.5635 | -105.5122
13.5671 | -105.5045
13.5708 | -105.4943
13.5745 | -105.4847
13.5781 | -105.4769
13.5818 | -105.4691
13.5855 | -105.4606
13.5891 | -105.4492
13.5928 | -105.4373
13.5964 | -105.4278
13.6001 | -105.4174
13.6038 | -105.4064
13.6074 | -105.3938
13.6111 | -105.3857
13.6148 | -105.3785
13.6184 | -105.3702
13.6221 | -105.3671
13.6258 | -105.3655
13.6294 | -105.3633
13.6331 | -105.3639
13.6367 | -105.3666
13.6404 | -105.3722
13.6441 | -105.3803
13.6477 | -105.3913
13.6514 | -105.4052
13.6551 | -105.4191
13.6587 | -105.429
13.6624 | -105.4456
13.666 | -105.4677
13.6697 | -105.4888
13.6734 | -105.5076
13.677 | -105.526
13.6807 | -105.5446
13.6844 | -105.5677
13.688 | -105.5894
13.6917 | -105.6084
13.6953 | -105.6283
13.699 | -105.6514
13.7027 | -105.6733
13.7063 | -105.6903
13.71 | -105.7106
13.7137 | -105.7305
13.7173 | -105.7503
13.721 | -105.7703
13.7246 | -105.7883
13.7283 | -105.8042
13.732 | -105.8203
13.7356 | -105.8255
13.7393 | -105.837
13.743 | -105.8487
13.7466 | -105.8601
13.7503 | -105.8683
13.7539 | -105.8716
13.7576 | -105.8739
13.7613 | -105.8735
13.7649 | -105.8683
13.7686 | -105.8604
13.7723 | -105.8492
13.7759 | -105.8354
13.7796 | -105.819
13.7833 | -105.805
13.7869 | -105.785
13.7906 | -105.7629
13.7942 | -105.7386
13.7979 | -105.7133
13.8016 | -105.6886
13.8052 | -105.6648
13.8089 | -105.6321
13.8126 | -105.5982
13.8162 | -105.5687
13.8199 | -105.5369
13.8235 | -105.5048
13.8272 | -105.4702
13.8309 | -105.4372
13.8345 | -105.4063
13.8382 | -105.3743
13.8419 | -105.3421
13.8455 | -105.3142
13.8492 | -105.2898
13.8528 | -105.267
13.8565 | -105.2458
13.8602 | -105.2257
13.8638 | -105.2095
13.8675 | -105.1995
13.8712 | -105.191
13.8748 | -105.1864
13.8785 | -105.1863
13.8821 | -105.1881
13.8858 | -105.1958
13.8895 | -105.2077
13.8931 | -105.2236
13.8968 | -105.2402
13.9005 | -105.2607
13.9041 | -105.2877
13.9078 | -105.3146
13.9115 | -105.3431
13.9151 | -105.3734
13.9188 | -105.4015
13.9224 | -105.4302
13.9261 | -105.4594
13.9298 | -105.4912
13.9334 | -105.524
13.9371 | -105.5543
13.9408 | -105.5811
13.9444 | -105.6058
13.9481 | -105.6279
13.9517 | -105.649
13.9554 | -105.6671
13.9591 | -105.6792
13.9627 | -105.6868
13.9664 | -105.691
13.9701 | -105.6941
13.9737 | -105.6966
13.9774 | -105.6931
13.981 | -105.6866
13.9847 | -105.673
13.9884 | -105.6612
13.992 | -105.6475
13.9957 | -105.6295
13.9994 | -105.6108
14.003 | -105.5894
14.0067 | -105.5672
14.0103 | -105.5495
14.014 | -105.5314
14.0177 | -105.5108
14.0213 | -105.4907
14.025 | -105.4713
14.0287 | -105.4558
14.0323 | -105.4405
14.036 | -105.4272
14.0396 | -105.4176
14.0433 | -105.412
14.047 | -105.4078
14.0506 | -105.4042
14.0543 | -105.4075
14.058 | -105.4127
14.0616 | -105.4188
14.0653 | -105.4257
14.069 | -105.434
14.0726 | -105.4472
14.0763 | -105.4609
14.0799 | -105.4738
14.0836 | -105.4898
14.0873 | -105.5092
14.0909 | -105.5289
14.0946 | -105.5461
14.0983 | -105.5681
14.1019 | -105.5884
14.1056 | -105.6073
14.1092 | -105.6294
14.1129 | -105.6505
14.1166 | -105.6712
14.1202 | -105.6937
14.1239 | -105.715
14.1276 | -105.7348
14.1312 | -105.7489
14.1349 | -105.7602
14.1385 | -105.7735
14.1422 | -105.7871
14.1459 | -105.8002
14.1495 | -105.8133
14.1532 | -105.8247
14.1569 | -105.8389
14.1605 | -105.8564
14.1642 | -105.8718
14.1678 | -105.8902
14.1715 | -105.9065
14.1752 | -105.9179
14.1788 | -105.9321
14.1825 | -105.948
14.1862 | -105.9648
14.1898 | -105.9834
14.1935 | -105.9988
14.1972 | -106.0137
14.2008 | -106.0301
14.2045 | -106.0446
14.2081 | -106.0573
14.2118 | -106.0696
14.2155 | -106.0791
14.2191 | -106.085
14.2228 | -106.0889
14.2265 | -106.0933
14.2301 | -106.0984
14.2338 | -106.1008
14.2374 | -106.1032
14.2411 | -106.1019
14.2448 | -106.0998
14.2484 | -106.0978
14.2521 | -106.0978
14.2558 | -106.0968
14.2594 | -106.0949
14.2631 | -106.0896
14.2667 | -106.0887
14.2704 | -106.0864
14.2741 | -106.0863
14.2777 | -106.0854
14.2814 | -106.0855
14.2851 | -106.0863
14.2887 | -106.0865
14.2924 | -106.0873
14.296 | -106.0886
14.2997 | -106.0935
14.3034 | -106.098
14.307 | -106.1017
14.3107 | -106.1066
14.3144 | -106.1081
14.318 | -106.1129
14.3217 | -106.1174
14.3254 | -106.1237
14.329 | -106.1293
14.3327 | -106.1327
14.3363 | -106.1411
14.34 | -106.1506
14.3437 | -106.1582
14.3473 | -106.1663
14.351 | -106.1759
14.3547 | -106.1847
14.3583 | -106.1925
14.362 | -106.1997
14.3656 | -106.2062
14.3693 | -106.2133
14.373 | -106.2196
14.3766 | -106.2255
14.3803 | -106.2342
14.384 | -106.2419
14.3876 | -106.248
14.3913 | -106.2538
14.3949 | -106.2586
14.3986 | -106.2666
14.4023 | -106.2732
14.4059 | -106.2765
14.4096 | -106.2804
14.4133 | -106.2846
14.4169 | -106.2866
14.4206 | -106.2905
14.4242 | -106.297
14.4279 | -106.3022
14.4316 | -106.3042
14.4352 | -106.3029
14.4389 | -106.3049
14.4426 | -106.3067
14.4462 | -106.3106
14.4499 | -106.3126
14.4536 | -106.3118
14.4572 | -106.3111
14.4609 | -106.3126
14.4645 | -106.3171
14.4682 | -106.3228
14.4719 | -106.3221
14.4755 | -106.3233
14.4792 | -106.3254
14.4829 | -106.3275
14.4865 | -106.3292
14.4902 | -106.3287
14.4938 | -106.3331
14.4975 | -106.3349
14.5012 | -106.3319
14.5048 | -106.3312
14.5085 | -106.3298
14.5122 | -106.3272
14.5158 | -106.3269
14.5195 | -106.3254
14.5231 | -106.3241
14.5268 | -106.3187
14.5305 | -106.3149
14.5341 | -106.3085
14.5378 | -106.3037
14.5415 | -106.298
14.5451 | -106.2889
14.5488 | -106.2812
14.5524 | -106.2757
14.5561 | -106.2679
14.5598 | -106.2574
14.5634 | -106.2467
14.5671 | -106.2351
14.5708 | -106.2233
14.5744 | -106.2112
14.5781 | -106.2016
14.5817 | -106.1876
14.5854 | -106.1755
14.5891 | -106.1627
14.5927 | -106.1474
14.5964 | -106.1327
14.6001 | -106.1203
14.6037 | -106.1096
14.6074 | -106.0961
14.6111 | -106.085
14.6147 | -106.076
14.6184 | -106.0717
14.622 | -106.0693
14.6257 | -106.0673
14.6294 | -106.0644
14.633 | -106.0596
14.6367 | -106.0581
14.6404 | -106.0597
14.644 | -106.0628
14.6477 | -106.0654
14.6513 | -106.0662
14.655 | -106.0694
14.6587 | -106.0754
14.6623 | -106.0848
14.666 | -106.0923
14.6697 | -106.0958
14.6733 | -106.1015
14.677 | -106.1128
14.6806 | -106.1268
14.6843 | -106.1387
14.688 | -106.1525
14.6916 | -106.1653
14.6953 | -106.1794
14.699 | -106.1934
14.7026 | -106.203
14.7063 | -106.213
14.7099 | -106.2269
14.7136 | -106.2399
14.7173 | -106.2501
14.7209 | -106.2563
14.7246 | -106.2633
14.7283 | -106.2724
14.7319 | -106.2789
14.7356 | -106.2816
14.7393 | -106.2806
14.7429 | -106.2772
14.7466 | -106.2673
14.7502 | -106.2578
14.7539 | -106.246
14.7576 | -106.2301
14.7612 | -106.2102
14.7649 | -106.1908
14.7686 | -106.1683
14.7722 | -106.1444
14.7759 | -106.121
14.7795 | -106.0962
14.7832 | -106.0702
14.7869 | -106.0412
14.7905 | -106.0142
14.7942 | -105.9842
14.7979 | -105.9591
14.8015 | -105.9319
14.8052 | -105.9068
14.8088 | -105.882
14.8125 | -105.8576
14.8162 | -105.8347
14.8198 | -105.8143
14.8235 | -105.7949
14.8272 | -105.7776
14.8308 | -105.7619
14.8345 | -105.7495
14.8381 | -105.7401
14.8418 | -105.7327
14.8455 | -105.7293
14.8491 | -105.7309
14.8528 | -105.7337
14.8565 | -105.7385
14.8601 | -105.7447
14.8638 | -105.7547
14.8675 | -105.7665
14.8711 | -105.7791
14.8748 | -105.7921
14.8784 | -105.807
14.8821 | -105.825
14.8858 | -105.8412
14.8894 | -105.8549
14.8931 | -105.8687
14.8968 | -105.8846
14.9004 | -105.8983
14.9041 | -105.9106
14.9077 | -105.9205
14.9114 | -105.9319
14.9151 | -105.9421
14.9187 | -105.9503
14.9224 | -105.9573
14.9261 | -105.9608
14.9297 | -105.9627
14.9334 | -105.9627
14.937 | -105.959
14.9407 | -105.9548
14.9444 | -105.9504
14.948 | -105.9471
14.9517 | -105.9411
14.9554 | -105.9356
14.959 | -105.928
14.9627 | -105.919
14.9663 | -105.9088
14.97 | -105.8963
14.9737 | -105.8829
14.9773 | -105.8689
14.981 | -105.8554
14.9847 | -105.8379
14.9883 | -105.8189
14.992 | -105.8031
14.9956 | -105.7874
14.9993 | -105.772
15.003 | -105.756
15.0066 | -105.7407
15.0103 | -105.7271
15.014 | -105.7154
15.0176 | -105.7031
15.0213 | -105.6897
15.025 | -105.6785
15.0286 | -105.6701
15.0323 | -105.6621
15.0359 | -105.6551
15.0396 | -105.6492
15.0433 | -105.6434
15.0469 | -105.6391
15.0506 | -105.6368
15.0543 | -105.635
15.0579 | -105.6343
15.0616 | -105.6362
15.0652 | -105.6361
15.0689 | -105.6386
15.0726 | -105.6434
15.0762 | -105.6494
15.0799 | -105.6548
15.0836 | -105.6622
15.0872 | -105.6707
15.0909 | -105.6818
15.0945 | -105.6956
15.0982 | -105.7103
15.1019 | -105.7245
15.1055 | -105.7405
15.1092 | -105.7573
15.1129 | -105.7736
15.1165 | -105.7917
15.1202 | -105.8125
15.1238 | -105.8357
15.1275 | -105.8583
15.1312 | -105.8819
15.1348 | -105.9079
15.1385 | -105.9327
15.1422 | -105.9575
15.1458 | -105.9856
15.1495 | -106.0126
15.1532 | -106.0399
15.1568 | -106.063
15.1605 | -106.085
15.1641 | -106.103
15.1678 | -106.1222
15.1715 | -106.1384
15.1751 | -106.1548
15.1788 | -106.1714
15.1825 | -106.1854
15.1861 | -106.1997
15.1898 | -106.2124
15.1934 | -106.2227
15.1971 | -106.2322
15.2008 | -106.2406
15.2044 | -106.247
15.2081 | -106.2491
15.2118 | -106.2481
15.2154 | -106.2464
15.2191 | -106.2442
15.2227 | -106.2402
15.2264 | -106.2324
15.2301 | -106.2215
15.2337 | -106.2113
15.2374 | -106.2009
15.2411 | -106.1885
15.2447 | -106.1732
15.2484 | -106.1583
15.252 | -106.1447
15.2557 | -106.1305
15.2594 | -106.1138
15.263 | -106.0924
15.2667 | -106.0701
15.2704 | -106.0505
15.274 | -106.0321
15.2777 | -106.0107
15.2814 | -105.989
15.285 | -105.9684
15.2887 | -105.9476
15.2923 | -105.9259
15.296 | -105.9071
15.2997 | -105.8874
15.3033 | -105.8667
15.307 | -105.8465
15.3107 | -105.8286
15.3143 | -105.8139
15.318 | -105.801
15.3216 | -105.7957
15.3253 | -105.786
15.329 | -105.7784
15.3326 | -105.7732
15.3363 | -105.7683
15.34 | -105.7614
15.3436 | -105.7569
15.3473 | -105.7565
15.3509 | -105.7578
15.3546 | -105.7594
15.3583 | -105.7627
15.3619 | -105.7679
15.3656 | -105.774
15.3693 | -105.7824
15.3729 | -105.7893
15.3766 | -105.7976
15.3802 | -105.8064
15.3839 | -105.8154
15.3876 | -105.8248
15.3912 | -105.8342
15.3949 | -105.8434
15.3986 | -105.8522
15.4022 | -105.8604
15.4059 | -105.8687
15.4096 | -105.875
15.4132 | -105.8821
15.4169 | -105.8904
15.4205 | -105.8992
15.4242 | -105.9062
15.4279 | -105.9135
15.4315 | -105.9239
15.4352 | -105.9337
15.4389 | -105.9412
15.4425 | -105.9461
15.4462 | -105.9511
15.4498 | -105.9582
15.4535 | -105.9647
15.4572 | -105.9685
15.4608 | -105.9688
15.4645 | -105.9686
15.4682 | -105.9671
15.4718 | -105.9662
15.4755 | -105.9652
15.4791 | -105.9632
15.4828 | -105.9614
15.4865 | -105.9587
15.4901 | -105.9532
15.4938 | -105.947
15.4975 | -105.9425
15.5011 | -105.9369
15.5048 | -105.929
15.5084 | -105.9238
15.5121 | -105.9183
15.5158 | -105.9111
15.5194 | -105.9042
15.5231 | -105.8983
15.5268 | -105.8943
15.5304 | -105.8895
15.5341 | -105.8823
15.5378 | -105.8783
15.5414 | -105.8754
15.5451 | -105.8716
15.5487 | -105.866
15.5524 | -105.8645
15.5561 | -105.8626
15.5597 | -105.8614
15.5634 | -105.8636
15.5671 | -105.8628
15.5707 | -105.8598
15.5744 | -105.8578
15.578 | -105.8582
15.5817 | -105.8597
15.5854 | -105.8622
15.589 | -105.8654
15.5927 | -105.8708
15.5964 | -105.8741
15.6 | -105.8774
15.6037 | -105.8811
15.6073 | -105.8868
15.611 | -105.8891
15.6147 | -105.8928
15.6183 | -105.8969
15.622 | -105.9008
15.6257 | -105.9068
15.6293 | -105.9126
15.633 | -105.9197
15.6366 | -105.9245
15.6403 | -105.9309
15.644 | -105.9362
15.6476 | -105.9392
15.6513 | -105.9427
15.655 | -105.9443
15.6586 | -105.9449
15.6623 | -105.9481
15.6659 | -105.9531
15.6696 | -105.9587
15.6733 | -105.9634
15.6769 | -105.968
15.6806 | -105.9749
15.6843 | -105.9804
15.6879 | -105.9836
15.6916 | -105.9865
15.6953 | -105.9867
15.6989 | -105.9876
15.7026 | -105.9898
15.7062 | -105.9917
15.7099 | -105.9954
15.7136 | -106.0001
15.7172 | -106.003
15.7209 | -106.0049
15.7246 | -106.007
15.7282 | -106.01
15.7319 | -106.0112
15.7355 | -106.014
15.7392 | -106.0171
15.7429 | -106.0214
15.7465 | -106.0252
15.7502 | -106.0252
15.7539 | -106.0256
15.7575 | -106.0294
15.7612 | -106.0342
15.7648 | -106.038
15.7685 | -106.0403
15.7722 | -106.0447
15.7758 | -106.048
15.7795 | -106.0533
15.7832 | -106.0599
15.7868 | -106.0688
15.7905 | -106.0818
15.7941 | -106.0929
15.7978 | -106.102
15.8015 | -106.1124
15.8051 | -106.1229
15.8088 | -106.1315
15.8125 | -106.1379
15.8161 | -106.1455
15.8198 | -106.1566
15.8235 | -106.1651
15.8271 | -106.1725
15.8308 | -106.1815
15.8344 | -106.1907
15.8381 | -106.1967
15.8418 | -106.1998
15.8454 | -106.202
15.8491 | -106.2014
15.8528 | -106.1972
15.8564 | -106.1962
15.8601 | -106.1926
15.8637 | -106.1921
15.8674 | -106.1898
15.8711 | -106.19
15.8747 | -106.1917
15.8784 | -106.1912
15.8821 | -106.1856
15.8857 | -106.1805
15.8894 | -106.1767
15.893 | -106.1711
15.8967 | -106.1613
15.9004 | -106.1533
15.904 | -106.1469
15.9077 | -106.1405
15.9114 | -106.1347
15.915 | -106.1296
15.9187 | -106.1242
15.9223 | -106.1208
15.926 | -106.1146
15.9297 | -106.1085
15.9333 | -106.1023
15.937 | -106.0925
15.9407 | -106.0832
15.9443 | -106.0773
15.948 | -106.0711
15.9517 | -106.0652
15.9553 | -106.0576
15.959 | -106.0493
15.9626 | -106.0439
15.9663 | -106.0396
15.97 | -106.0345
15.9736 | -106.0308
15.9773 | -106.0302
15.981 | -106.0304
15.9846 | -106.0261
15.9883 | -106.0218
15.9919 | -106.0211
15.9956 | -106.0235
15.9993 | -106.0258
16.0029 | -106.0282
16.0066 | -106.0304
16.0103 | -106.0328
16.0139 | -106.0349
16.0176 | -106.0344
16.0212 | -106.0361
16.0249 | -106.0413
16.0286 | -106.0458
16.0322 | -106.0489
16.0359 | -106.0524
16.0396 | -106.0575
16.0432 | -106.0625
16.0469 | -106.0656
16.0505 | -106.0712
16.0542 | -106.0768
16.0579 | -106.0823
16.0615 | -106.0888
16.0652 | -106.0954
16.0689 | -106.1004
16.0725 | -106.1032
16.0762 | -106.1105
16.0799 | -106.1166
16.0835 | -106.1221
16.0872 | -106.126
16.0908 | -106.1311
16.0945 | -106.1333
16.0982 | -106.1344
16.1018 | -106.1335
16.1055 | -106.1307
16.1092 | -106.1288
16.1128 | -106.1316
16.1165 | -106.1319
16.1201 | -106.1281
16.1238 | -106.1247
16.1275 | -106.1191
16.1311 | -106.1116
16.1348 | -106.1031
16.1385 | -106.0895
16.1421 | -106.0763
16.1458 | -106.064
16.1494 | -106.0504
16.1531 | -106.035
16.1568 | -106.0201
16.1604 | -106.0034
16.1641 | -105.9859
16.1678 | -105.9707
16.1714 | -105.955
16.1751 | -105.9359
16.1787 | -105.9166
16.1824 | -105.9001
16.1861 | -105.8859
16.1897 | -105.8735
16.1934 | -105.8581
16.1971 | -105.8425
16.2007 | -105.8278
16.2044 | -105.8143
16.208 | -105.803
16.2117 | -105.7915
16.2154 | -105.785
16.219 | -105.7791
16.2227 | -105.7739
16.2264 | -105.7701
16.23 | -105.7678
16.2337 | -105.766
16.2374 | -105.7653
16.241 | -105.767
16.2447 | -105.7728
16.2483 | -105.7828
16.252 | -105.7916
16.2557 | -105.8
16.2593 | -105.8117
16.263 | -105.8236
16.2667 | -105.8378
16.2703 | -105.855
16.274 | -105.8712
16.2776 | -105.8892
16.2813 | -105.9081
16.285 | -105.928
16.2886 | -105.9464
16.2923 | -105.9661
16.296 | -105.9879
16.2996 | -106.0112
16.3033 | -106.0351
16.3069 | -106.0559
16.3106 | -106.0786
16.3143 | -106.1019
16.3179 | -106.1261
16.3216 | -106.1478
16.3253 | -106.1672
16.3289 | -106.1863
16.3326 | -106.2073
16.3362 | -106.2291
16.3399 | -106.2465
16.3436 | -106.2612
16.3472 | -106.2788
16.3509 | -106.2943
16.3546 | -106.3071
16.3582 | -106.3197
16.3619 | -106.3287
16.3656 | -106.3377
16.3692 | -106.3482
16.3729 | -106.3552
16.3765 | -106.3587
16.3802 | -106.358
16.3839 | -106.362
16.3875 | -106.3651
16.3912 | -106.3686
16.3949 | -106.3708
16.3985 | -106.3715
16.4022 | -106.3713
16.4058 | -106.3686
16.4095 | -106.3672
16.4132 | -106.3668
16.4168 | -106.3643
16.4205 | -106.3624
16.4242 | -106.3608
16.4278 | -106.3587
16.4315 | -106.3581
16.4351 | -106.3572
16.4388 | -106.3565
16.4425 | -106.3549
16.4461 | -106.3573
16.4498 | -106.3604
16.4535 | -106.364
16.4571 | -106.3685
16.4608 | -106.3726
16.4644 | -106.3772
16.4681 | -106.3831
16.4718 | -106.3915
16.4754 | -106.4022
16.4791 | -106.4102
16.4828 | -106.4196
16.4864 | -106.4322
16.4901 | -106.4457
16.4937 | -106.4585
16.4974 | -106.4752
16.5011 | -106.4928
16.5047 | -106.5073
16.5084 | -106.5207
16.5121 | -106.5341
16.5157 | -106.5483
16.5194 | -106.5632
16.5231 | -106.578
16.5267 | -106.5908
16.5304 | -106.6021
16.534 | -106.6122
16.5377 | -106.6212
16.5414 | -106.6304
16.545 | -106.6389
16.5487 | -106.6463
16.5524 | -106.6517
16.556 | -106.6578
16.5597 | -106.6643
16.5633 | -106.6647
16.567 | -106.662
16.5707 | -106.6584
16.5743 | -106.655
16.578 | -106.6525
16.5817 | -106.647
16.5853 | -106.6442
16.589 | -106.6404
16.5926 | -106.6311
16.5963 | -106.6199
16.6 | -106.6113
16.6036 | -106.6
16.6073 | -106.5873
16.611 | -106.5767
16.6146 | -106.5655
16.6183 | -106.5521
16.6219 | -106.5369
16.6256 | -106.5211
16.6293 | -106.5065
16.6329 | -106.4929
16.6366 | -106.4797
16.6403 | -106.4647
16.6439 | -106.4522
16.6476 | -106.4374
16.6513 | -106.421
16.6549 | -106.4095
16.6586 | -106.3971
16.6622 | -106.3813
16.6659 | -106.3671
16.6696 | -106.3533
16.6732 | -106.3406
16.6769 | -106.3288
16.6806 | -106.3175
16.6842 | -106.3042
16.6879 | -106.293
16.6915 | -106.2831
16.6952 | -106.272
16.6989 | -106.2636
16.7025 | -106.2575
16.7062 | -106.251
16.7099 | -106.2452
16.7135 | -106.2394
16.7172 | -106.2369
16.7208 | -106.2336
16.7245 | -106.2314
16.7282 | -106.2305
16.7318 | -106.2312
16.7355 | -106.2355
16.7392 | -106.24
16.7428 | -106.2457
16.7465 | -106.2535
16.7501 | -106.2619
16.7538 | -106.271
16.7575 | -106.2814
16.7611 | -106.2942
16.7648 | -106.3087
16.7685 | -106.3252
16.7721 | -106.3429
16.7758 | -106.3609
16.7795 | -106.3789
16.7831 | -106.3959
16.7868 | -106.4141
16.7904 | -106.4324
16.7941 | -106.451
16.7978 | -106.468
16.8014 | -106.4866
16.8051 | -106.5076
16.8088 | -106.5267
16.8124 | -106.5451
16.8161 | -106.5597
16.8197 | -106.5717
16.8234 | -106.5804
16.8271 | -106.5899
16.8307 | -106.6041
16.8344 | -106.6183
16.8381 | -106.6239
16.8417 | -106.6285
16.8454 | -106.6328
16.849 | -106.6328
16.8527 | -106.632
16.8564 | -106.6309
16.86 | -106.6298
16.8637 | -106.6267
16.8674 | -106.6216
16.871 | -106.6169
16.8747 | -106.6093
16.8783 | -106.6029
16.882 | -106.5974
16.8857 | -106.5923
16.8893 | -106.5869
16.893 | -106.5797
16.8967 | -106.5725
16.9003 | -106.5687
16.904 | -106.5664
16.9077 | -106.5634
16.9113 | -106.5608
16.915 | -106.5577
16.9186 | -106.5541
16.9223 | -106.549
16.926 | -106.5433
16.9296 | -106.5408
16.9333 | -106.5403
16.937 | -106.5409
16.9406 | -106.5409
16.9443 | -106.5406
16.9479 | -106.5409
16.9516 | -106.5423
16.9553 | -106.5447
16.9589 | -106.5453
16.9626 | -106.5481
16.9663 | -106.5507
16.9699 | -106.5521
16.9736 | -106.5511
16.9772 | -106.5502
16.9809 | -106.5488
16.9846 | -106.5462
16.9882 | -106.5456
16.9919 | -106.545
16.9956 | -106.5425
16.9992 | -106.5403
17.0029 | -106.5375
17.0065 | -106.5355
17.0102 | -106.532
17.0139 | -106.5299
17.0175 | -106.5288
17.0212 | -106.527
17.0249 | -106.5246
17.0285 | -106.5211
17.0322 | -106.5193
17.0358 | -106.5181
17.0395 | -106.5155
17.0432 | -106.5142
17.0468 | -106.5111
17.0505 | -106.5103
17.0542 | -106.5092
17.0578 | -106.5098
17.0615 | -106.5112
17.0652 | -106.5127
17.0688 | -106.5158
17.0725 | -106.517
17.0761 | -106.5162
17.0798 | -106.517
17.0835 | -106.5193
17.0871 | -106.5224
17.0908 | -106.5232
17.0945 | -106.5238
17.0981 | -106.5244
17.1018 | -106.5264
17.1054 | -106.529
17.1091 | -106.5293
17.1128 | -106.5281
17.1164 | -106.5302
17.1201 | -106.5308
17.1238 | -106.5318
17.1274 | -106.5343
17.1311 | -106.5368
17.1347 | -106.5361
17.1384 | -106.5353
17.1421 | -106.5373
17.1457 | -106.5378
17.1494 | -106.538
17.1531 | -106.5387
17.1567 | -106.5364
17.1604 | -106.5342
17.164 | -106.5296
17.1677 | -106.5259
17.1714 | -106.5218
17.175 | -106.5151
17.1787 | -106.5067
17.1824 | -106.5004
17.186 | -106.4928
17.1897 | -106.4848
17.1934 | -106.4779
17.197 | -106.4696
17.2007 | -106.4607
17.2043 | -106.4547
17.208 | -106.4512
17.2117 | -106.446
17.2153 | -106.442
17.219 | -106.4381
17.2227 | -106.4326
17.2263 | -106.4301
17.23 | -106.4259
17.2336 | -106.4205
17.2373 | -106.4172
17.241 | -106.4151
17.2446 | -106.4122
17.2483 | -106.4104
17.252 | -106.4085
17.2556 | -106.408
17.2593 | -106.406
17.2629 | -106.4013
17.2666 | -106.3974
17.2703 | -106.3958
17.2739 | -106.3939
17.2776 | -106.3915
17.2813 | -106.3885
17.2849 | -106.3852
17.2886 | -106.3821
17.2922 | -106.3762
17.2959 | -106.3699
17.2996 | -106.366
17.3032 | -106.3623
17.3069 | -106.359
17.3106 | -106.3553
17.3142 | -106.3502
17.3179 | -106.3469
17.3216 | -106.342
17.3252 | -106.3387
17.3289 | -106.337
17.3325 | -106.3332
17.3362 | -106.3298
17.3399 | -106.327
17.3435 | -106.3254
17.3472 | -106.3239
17.3509 | -106.3231
17.3545 | -106.3257
17.3582 | -106.3284
17.3618 | -106.3303
17.3655 | -106.3314
17.3692 | -106.3317
17.3728 | -106.334
17.3765 | -106.3369
17.3802 | -106.3413
17.3838 | -106.3482
17.3875 | -106.3525
17.3911 | -106.3556
17.3948 | -106.3617
17.3985 | -106.3675
17.4021 | -106.3714
17.4058 | -106.3764
17.4095 | -106.3828
17.4131 | -106.3882
17.4168 | -106.3942
17.4204 | -106.3966
17.4241 | -106.3997
17.4278 | -106.4012
17.4314 | -106.4054
17.4351 | -106.4062
17.4388 | -106.4081
17.4424 | -106.4092
17.4461 | -106.4117
17.4498 | -106.413
17.4534 | -106.4112
17.4571 | -106.4094
17.4607 | -106.4103
17.4644 | -106.4102
17.4681 | -106.4127
17.4717 | -106.4152
17.4754 | -106.4147
17.4791 | -106.4188
17.4827 | -106.4209
17.4864 | -106.4206
17.49 | -106.4202
17.4937 | -106.4175
17.4974 | -106.4157
17.501 | -106.4166
17.5047 | -106.4169
17.5084 | -106.4175
17.512 | -106.4178
17.5157 | -106.4158
17.5193 | -106.4143
17.523 | -106.4137
17.5267 | -106.4124
17.5303 | -106.4109
17.534 | -106.4104
17.5377 | -106.4055
17.5413 | -106.4044
17.545 | -106.4028
17.5486 | -106.3994
17.5523 | -106.3947
17.556 | -106.3915
17.5596 | -106.3903
17.5633 | -106.3874
17.567 | -106.3845
17.5706 | -106.3851
17.5743 | -106.3828
17.578 | -106.3782
17.5816 | -106.3743
17.5853 | -106.3707
17.5889 | -106.366
17.5926 | -106.3622
17.5963 | -106.3597
17.5999 | -106.3565
17.6036 | -106.3512
17.6073 | -106.3467
17.6109 | -106.341
17.6146 | -106.3364
17.6182 | -106.3314
17.6219 | -106.3262
17.6256 | -106.3222
17.6292 | -106.3192
17.6329 | -106.315
17.6366 | -106.3075
17.6402 | -106.3039
17.6439 | -106.3007
17.6475 | -106.2986
17.6512 | -106.2966
17.6549 | -106.2933
17.6585 | -106.2892
17.6622 | -106.2868
17.6659 | -106.2829
17.6695 | -106.2804
17.6732 | -106.2787
17.6768 | -106.2766
17.6805 | -106.2739
17.6842 | -106.2718
17.6878 | -106.2715
17.6915 | -106.2701
17.6952 | -106.2675
17.6988 | -106.2657
17.7025 | -106.2638
17.7061 | -106.2618
17.7098 | -106.2583
17.7135 | -106.2564
17.7171 | -106.2567
17.7208 | -106.2543
17.7245 | -106.2526
17.7281 | -106.2532
17.7318 | -106.2505
17.7355 | -106.2513
17.7391 | -106.2534
17.7428 | -106.2548
17.7464 | -106.2582
17.7501 | -106.2615
17.7538 | -106.2642
17.7574 | -106.2659
17.7611 | -106.27
17.7648 | -106.2764
17.7684 | -106.2818
17.7721 | -106.285
17.7757 | -106.2914
17.7794 | -106.2966
17.7831 | -106.3034
17.7867 | -106.3104
17.7904 | -106.3188
17.7941 | -106.3252
17.7977 | -106.3315
17.8014 | -106.3376
17.805 | -106.3445
17.8087 | -106.3489
17.8124 | -106.3514
17.816 | -106.358
17.8197 | -106.3636
17.8234 | -106.3668
17.827 | -106.3693
17.8307 | -106.371
17.8343 | -106.3702
17.838 | -106.3703
17.8417 | -106.3702
17.8453 | -106.3654
17.849 | -106.3594
17.8527 | -106.3541
17.8563 | -106.35
17.86 | -106.3454
17.8637 | -106.3396
17.8673 | -106.3326
17.871 | -106.3257
17.8746 | -106.3194
17.8783 | -106.3108
17.882 | -106.3024
17.8856 | -106.2937
17.8893 | -106.2883
17.893 | -106.2805
17.8966 | -106.2716
17.9003 | -106.2649
17.9039 | -106.2579
17.9076 | -106.2521
17.9113 | -106.2491
17.9149 | -106.244
17.9186 | -106.2422
17.9223 | -106.2418
17.9259 | -106.2401
17.9296 | -106.2406
17.9332 | -106.2422
17.9369 | -106.2434
17.9406 | -106.2461
17.9442 | -106.2505
17.9479 | -106.2573
17.9516 | -106.263
17.9552 | -106.2694
17.9589 | -106.2798
17.9625 | -106.2898
17.9662 | -106.3014
17.9699 | -106.3099
17.9735 | -106.3182
17.9772 | -106.3264
17.9809 | -106.3319
17.9845 | -106.3376
17.9882 | -106.3478
17.9918 | -106.3556
17.9955 | -106.358
17.9992 | -106.3616
18.0028 | -106.3644
18.0065 | -106.3664
18.0102 | -106.3648
18.0138 | -106.3628
18.0175 | -106.3603
18.0212 | -106.3574
18.0248 | -106.3546
18.0285 | -106.3495
18.0321 | -106.3401
18.0358 | -106.332
18.0395 | -106.3221
18.0431 | -106.3096
18.0468 | -106.2961
18.0505 | -106.2823
18.0541 | -106.2673
18.0578 | -106.2518
18.0614 | -106.2374
18.0651 | -106.2225
18.0688 | -106.2057
18.0724 | -106.1885
18.0761 | -106.1696
18.0798 | -106.1515
18.0834 | -106.1367
18.0871 | -106.1236
18.0907 | -106.1072
18.0944 | -106.0913
18.0981 | -106.0772
18.1017 | -106.0668
18.1054 | -106.057
18.1091 | -106.0465
18.1127 | -106.0357
18.1164 | -106.0283
18.12 | -106.0239
18.1237 | -106.0216
18.1274 | -106.0199
18.131 | -106.0198
18.1347 | -106.0203
18.1384 | -106.0199
18.142 | -106.0229
18.1457 | -106.0268
18.1494 | -106.0321
18.153 | -106.0365
18.1567 | -106.042
18.1603 | -106.05
18.164 | -106.0576
18.1677 | -106.0626
18.1713 | -106.0708
18.175 | -106.0788
18.1787 | -106.0877
18.1823 | -106.0976
18.186 | -106.1058
18.1896 | -106.1127
18.1933 | -106.1212
18.197 | -106.1292
18.2006 | -106.1361
18.2043 | -106.141
18.208 | -106.1452
18.2116 | -106.1497
18.2153 | -106.1525
18.2189 | -106.1558
18.2226 | -106.1559
18.2263 | -106.1554
18.2299 | -106.155
18.2336 | -106.156
18.2373 | -106.1554
18.2409 | -106.1572
18.2446 | -106.156
18.2482 | -106.1499
18.2519 | -106.1436
18.2556 | -106.1412
18.2592 | -106.137
18.2629 | -106.1333
18.2666 | -106.1287
18.2702 | -106.1253
18.2739 | -106.1238
18.2776 | -106.1201
18.2812 | -106.1158
18.2849 | -106.1117
18.2885 | -106.1077
18.2922 | -106.1061
18.2959 | -106.1061
18.2995 | -106.1049
18.3032 | -106.1058
18.3069 | -106.1079
18.3105 | -106.1112
18.3142 | -106.1145
18.3178 | -106.1197
18.3215 | -106.1249
18.3252 | -106.1325
18.3288 | -106.1405
18.3325 | -106.1488
18.3362 | -106.1605
18.3398 | -106.1737
18.3435 | -106.1852
18.3471 | -106.1968
18.3508 | -106.213
18.3545 | -106.227
18.3581 | -106.2381
18.3618 | -106.2467
18.3655 | -106.2567
18.3691 | -106.2703
18.3728 | -106.2831
18.3764 | -106.293
18.3801 | -106.302
18.3838 | -106.3124
18.3874 | -106.3248
18.3911 | -106.3344
18.3948 | -106.3439
18.3984 | -106.3535
18.4021 | -106.3597
18.4058 | -106.3657
18.4094 | -106.3688
18.4131 | -106.3725
18.4167 | -106.3736
18.4204 | -106.3736
18.4241 | -106.3732
18.4277 | -106.3705
18.4314 | -106.3669
18.4351 | -106.3633
18.4387 | -106.3576
18.4424 | -106.3496
18.446 | -106.3436
18.4497 | -106.3362
18.4534 | -106.3282
18.457 | -106.3207
18.4607 | -106.3132
18.4644 | -106.3065
18.468 | -106.2994
18.4717 | -106.2917
18.4753 | -106.2834
18.479 | -106.2753
18.4827 | -106.2693
18.4863 | -106.2616
18.49 | -106.2539
18.4937 | -106.2495
18.4973 | -106.241
18.501 | -106.235
18.5046 | -106.2292
18.5083 | -106.2241
18.512 | -106.2218
18.5156 | -106.2209
18.5193 | -106.2195
18.523 | -106.2178
18.5266 | -106.2173
18.5303 | -106.2171
18.534 | -106.2174
18.5376 | -106.2172
18.5413 | -106.2198
18.5449 | -106.2248
18.5486 | -106.2316
18.5523 | -106.239
18.5559 | -106.2463
18.5596 | -106.2535
18.5633 | -106.2612
18.5669 | -106.2708
18.5706 | -106.2813
18.5742 | -106.2908
18.5779 | -106.3033
18.5816 | -106.3158
18.5852 | -106.3276
18.5889 | -106.3406
18.5926 | -106.3571
18.5962 | -106.3724
18.5999 | -106.3865
18.6035 | -106.4009
18.6072 | -106.4141
18.6109 | -106.4256
18.6145 | -106.439
18.6182 | -106.4536
18.6219 | -106.4682
18.6255 | -106.4836
18.6292 | -106.4967
18.6328 | -106.5059
18.6365 | -106.5158
18.6402 | -106.525
18.6438 | -106.5331
18.6475 | -106.5412
18.6512 | -106.5459
18.6548 | -106.5512
18.6585 | -106.5588
18.6621 | -106.5635
18.6658 | -106.5692
18.6695 | -106.5725
18.6731 | -106.575
18.6768 | -106.5746
18.6805 | -106.5731
18.6841 | -106.5714
18.6878 | -106.5685
18.6915 | -106.5673
18.6951 | -106.5627
18.6988 | -106.557
18.7024 | -106.5496
18.7061 | -106.5432
18.7098 | -106.5365
18.7134 | -106.5297
18.7171 | -106.5219
18.7208 | -106.513
18.7244 | -106.5051
18.7281 | -106.4954
18.7317 | -106.4873
18.7354 | -106.4807
18.7391 | -106.4736
18.7427 | -106.4655
18.7464 | -106.4567
18.7501 | -106.4509
18.7537 | -106.445
18.7574 | -106.4404
18.761 | -106.434
18.7647 | -106.4302
18.7684 | -106.4271
18.772 | -106.4238
18.7757 | -106.4212
18.7794 | -106.4174
18.783 | -106.4128
18.7867 | -106.4084
18.7903 | -106.4058
18.794 | -106.4042
18.7977 | -106.4025
18.8013 | -106.4017
18.805 | -106.4009
18.8087 | -106.396
18.8123 | -106.3899
18.816 | -106.3849
18.8197 | -106.3826
18.8233 | -106.3795
18.827 | -106.377
18.8306 | -106.3773
18.8343 | -106.3757
18.838 | -106.3745
18.8416 | -106.3731
18.8453 | -106.3704
18.849 | -106.3688
18.8526 | -106.3667
18.8563 | -106.366
18.8599 | -106.3652
18.8636 | -106.3649
18.8673 | -106.3642
18.8709 | -106.3647
18.8746 | -106.3667
18.8783 | -106.3695
18.8819 | -106.3728
18.8856 | -106.3766
18.8892 | -106.3782
18.8929 | -106.3815
18.8966 | -106.3853
18.9002 | -106.3914
18.9039 | -106.3987
18.9076 | -106.4055
18.9112 | -106.4133
18.9149 | -106.4195
18.9185 | -106.4244
18.9222 | -106.4295
18.9259 | -106.4358
18.9295 | -106.4433
18.9332 | -106.4482
18.9369 | -106.4531
18.9405 | -106.4586
18.9442 | -106.465
18.9478 | -106.4717
18.9515 | -106.477
18.9552 | -106.4819
18.9588 | -106.4843
18.9625 | -106.4846
18.9662 | -106.4846
18.9698 | -106.4862
18.9735 | -106.4877
18.9772 | -106.4879
18.9808 | -106.4846
18.9845 | -106.4819
18.9881 | -106.4809
18.9918 | -106.4777
18.9955 | -106.4763
18.9991 | -106.4742
19.0028 | -106.4702
19.0065 | -106.4663
19.0101 | -106.4629
19.0138 | -106.4603
19.0174 | -106.4576
19.0211 | -106.4546
19.0248 | -106.4515
19.0284 | -106.4471
19.0321 | -106.4422
19.0358 | -106.4365
19.0394 | -106.432
19.0431 | -106.4249
19.0467 | -106.4208
19.0504 | -106.4193
19.0541 | -106.4169
19.0577 | -106.4157
19.0614 | -106.412
19.0651 | -106.4048
19.0687 | -106.3995
19.0724 | -106.3952
19.076 | -106.3907
19.0797 | -106.3889
19.0834 | -106.3879
19.087 | -106.3848
19.0907 | -106.3809
19.0944 | -106.3773
19.098 | -106.3739
19.1017 | -106.3694
19.1054 | -106.3672
19.109 | -106.3655
19.1127 | -106.3643
19.1163 | -106.3625
19.12 | -106.3581
19.1237 | -106.3564
19.1273 | -106.3542
19.131 | -106.3512
19.1347 | -106.3467
19.1383 | -106.347
19.142 | -106.3463
19.1456 | -106.3458
19.1493 | -106.3461
19.153 | -106.345
19.1566 | -106.3443
19.1603 | -106.3472
19.164 | -106.348
19.1676 | -106.3483
19.1713 | -106.3485
19.1749 | -106.3503
19.1786 | -106.3539
19.1823 | -106.3576
19.1859 | -106.36
19.1896 | -106.3656
19.1933 | -106.3718
19.1969 | -106.3746
19.2006 | -106.3785
19.2042 | -106.3825
19.2079 | -106.3868
19.2116 | -106.3898
19.2152 | -106.3915
19.2189 | -106.3932
19.2226 | -106.3965
19.2262 | -106.3997
19.2299 | -106.4016
19.2336 | -106.4027
19.2372 | -106.4037
19.2409 | -106.4061
19.2445 | -106.407
19.2482 | -106.4062
19.2519 | -106.4089
19.2555 | -106.4097
19.2592 | -106.4082
19.2629 | -106.4094
19.2665 | -106.4099
19.2702 | -106.4102
19.2738 | -106.4108
19.2775 | -106.411
19.2812 | -106.4105
19.2848 | -106.4105
19.2885 | -106.4096
19.2922 | -106.4087
19.2958 | -106.4081
19.2995 | -106.4072
19.3031 | -106.4062
19.3068 | -106.4051
19.3105 | -106.4044
19.3141 | -106.4032
19.3178 | -106.4046
19.3215 | -106.4018
19.3251 | -106.3989
19.3288 | -106.3974
19.3324 | -106.3954
19.3361 | -106.3943
19.3398 | -106.3934
19.3434 | -106.392
19.3471 | -106.3924
19.3508 | -106.3937
19.3544 | -106.3934
19.3581 | -106.3949
19.3618 | -106.3983
19.3654 | -106.4008
19.3691 | -106.4027
19.3727 | -106.4059
19.3764 | -106.4092
19.3801 | -106.4106
19.3837 | -106.4123
19.3874 | -106.417
19.3911 | -106.4246
19.3947 | -106.4353
19.3984 | -106.4442
19.402 | -106.4511
19.4057 | -106.4584
19.4094 | -106.4687
19.413 | -106.4768
19.4167 | -106.4828
19.4204 | -106.4891
19.424 | -106.4948
19.4277 | -106.5027
19.4313 | -106.5097
19.435 | -106.5169
19.4387 | -106.5253
19.4423 | -106.5349
19.446 | -106.5438
19.4497 | -106.5511
19.4533 | -106.5579
19.457 | -106.5656
19.4606 | -106.5735
19.4643 | -106.5815
19.468 | -106.5884
19.4716 | -106.5976
19.4753 | -106.6056
19.479 | -106.6138
19.4826 | -106.6219
19.4863 | -106.6294
19.49 | -106.6355
19.4936 | -106.6405
19.4973 | -106.6439
19.5009 | -106.6473
19.5046 | -106.6527
19.5083 | -106.6584
19.5119 | -106.6617
19.5156 | -106.6627
19.5193 | -106.6653
19.5229 | -106.6663
19.5266 | -106.6643
19.5302 | -106.6608
19.5339 | -106.6574
19.5376 | -106.6525
19.5412 | -106.6481
19.5449 | -106.6439
19.5486 | -106.6377
19.5522 | -106.6325
19.5559 | -106.6283
19.5595 | -106.6249
19.5632 | -106.6196
19.5669 | -106.6144
19.5705 | -106.61
19.5742 | -106.6037
19.5779 | -106.5979
19.5815 | -106.5945
19.5852 | -106.591
19.5888 | -106.5866
19.5925 | -106.5817
19.5962 | -106.5759
19.5998 | -106.5701
19.6035 | -106.5645
19.6072 | -106.5602
19.6108 | -106.554
19.6145 | -106.5483
19.6181 | -106.5435
19.6218 | -106.5386
19.6255 | -106.5329
19.6291 | -106.5273
19.6328 | -106.5217
19.6365 | -106.5159
19.6401 | -106.5099
19.6438 | -106.5043
19.6475 | -106.5001
19.6511 | -106.4948
19.6548 | -106.4909
19.6584 | -106.4885
19.6621 | -106.4859
19.6658 | -106.4815
19.6694 | -106.4753
19.6731 | -106.4697
19.6768 | -106.4638
19.6804 | -106.4592
19.6841 | -106.4565
19.6877 | -106.4542
19.6914 | -106.452
19.6951 | -106.4523
19.6987 | -106.4524
19.7024 | -106.4518
19.7061 | -106.4486
19.7097 | -106.4478
19.7134 | -106.4465
19.717 | -106.4457
19.7207 | -106.4454
19.7244 | -106.4446
19.728 | -106.4419
19.7317 | -106.4404
19.7354 | -106.438
19.739 | -106.4369
19.7427 | -106.4376
19.7463 | -106.4379
19.75 | -106.4365
19.7537 | -106.4361
19.7573 | -106.4348
19.761 | -106.434
19.7647 | -106.434
19.7683 | -106.4344
19.772 | -106.4346
19.7757 | -106.4359
19.7793 | -106.4378
19.783 | -106.4374
19.7866 | -106.4359
19.7903 | -106.4361
19.794 | -106.4362
19.7976 | -106.4326
19.8013 | -106.4315
19.805 | -106.4315
19.8086 | -106.4322
19.8123 | -106.433
19.8159 | -106.4304
19.8196 | -106.4272
19.8233 | -106.4247
19.8269 | -106.4238
19.8306 | -106.422
19.8343 | -106.42
19.8379 | -106.4167
19.8416 | -106.414
19.8452 | -106.4114
19.8489 | -106.4083
19.8526 | -106.4044
19.8562 | -106.4014
19.8599 | -106.398
19.8636 | -106.3922
19.8672 | -106.3875
19.8709 | -106.383
19.8745 | -106.379
19.8782 | -106.3762
19.8819 | -106.3736
19.8855 | -106.3701
19.8892 | -106.3673
19.8929 | -106.3664
19.8965 | -106.3663
19.9002 | -106.3638
19.9039 | -106.3609
19.9075 | -106.3581
19.9112 | -106.3566
19.9148 | -106.3542
19.9185 | -106.3522
19.9222 | -106.3511
19.9258 | -106.3513
19.9295 | -106.3527
19.9332 | -106.3554
19.9368 | -106.3577
19.9405 | -106.3615
19.9441 | -106.363
19.9478 | -106.3643
19.9515 | -106.3682
19.9551 | -106.3711
19.9588 | -106.3732
19.9625 | -106.3753
19.9661 | -106.3772
19.9698 | -106.3782
19.9734 | -106.3803
19.9771 | -106.382
19.9808 | -106.3858
19.9844 | -106.3908
19.9881 | -106.3954
19.9918 | -106.3998
19.9954 | -106.4024
19.9991 | -106.4047
20.0027 | -106.4078
20.0064 | -106.4115
20.0101 | -106.4175
20.0137 | -106.4222
20.0174 | -106.4282
20.0211 | -106.434
20.0247 | -106.4366
20.0284 | -106.4414
20.0321 | -106.4473
20.0357 | -106.453
20.0394 | -106.4597
20.043 | -106.466
20.0467 | -106.47
20.0504 | -106.4733
20.054 | -106.4776
20.0577 | -106.4824
20.0614 | -106.4879
20.065 | -106.4942
20.0687 | -106.5002
20.0723 | -106.5064
20.076 | -106.5118
20.0797 | -106.5195
20.0833 | -106.5286
20.087 | -106.5373
20.0907 | -106.545
20.0943 | -106.5506
20.098 | -106.5554
20.1016 | -106.5641
20.1053 | -106.5727
20.109 | -106.5803
20.1126 | -106.5863
20.1163 | -106.5911
20.12 | -106.6
20.1236 | -106.6109
20.1273 | -106.6211
20.1309 | -106.6307
20.1346 | -106.6401
20.1383 | -106.6498
20.1419 | -106.6595
20.1456 | -106.6704
20.1493 | -106.683
20.1529 | -106.6941
20.1566 | -106.7026
20.1602 | -106.7111
20.1639 | -106.7204
20.1676 | -106.7309
20.1712 | -106.7387
20.1749 | -106.7439
20.1786 | -106.7473
20.1822 | -106.7505
20.1859 | -106.7557
20.1896 | -106.7598
20.1932 | -106.7619
20.1969 | -106.7634
20.2005 | -106.7661
20.2042 | -106.7657
20.2079 | -106.7646
20.2115 | -106.7625
20.2152 | -106.7593
20.2189 | -106.7536
20.2225 | -106.7488
20.2262 | -106.7424
20.2298 | -106.7388
20.2335 | -106.7358
20.2372 | -106.7302
20.2408 | -106.7247
20.2445 | -106.722
20.2482 | -106.7173
20.2518 | -106.7114
20.2555 | -106.7046
20.2591 | -106.6954
20.2628 | -106.6868
20.2665 | -106.6771
20.2701 | -106.6669
20.2738 | -106.6582
20.2775 | -106.6481
20.2811 | -106.6393
20.2848 | -106.6314
20.2884 | -106.6223
20.2921 | -106.6121
20.2958 | -106.6019
20.2994 | -106.5945
20.3031 | -106.5849
20.3068 | -106.5785
20.3104 | -106.5728
20.3141 | -106.5652
20.3178 | -106.5576
20.3214 | -106.5501
20.3251 | -106.5436
20.3287 | -106.5369
20.3324 | -106.5328
20.3361 | -106.5287
20.3397 | -106.525
20.3434 | -106.52
20.3471 | -106.5143
20.3507 | -106.5098
20.3544 | -106.507
20.358 | -106.5016
20.3617 | -106.4979
20.3654 | -106.4938
20.369 | -106.4895
20.3727 | -106.484
20.3764 | -106.4791
20.38 | -106.4755
20.3837 | -106.4735
20.3873 | -106.4702
20.391 | -106.4654
20.3947 | -106.4593
20.3983 | -106.4534
20.402 | -106.4497
20.4057 | -106.4474
20.4093 | -106.4429
20.413 | -106.4358
20.4166 | -106.4293
20.4203 | -106.4246
20.424 | -106.421
20.4276 | -106.4177
20.4313 | -106.416
20.435 | -106.4149
20.4386 | -106.4121
20.4423 | -106.4108
20.4459 | -106.4127
20.4496 | -106.4129
20.4533 | -106.4105
20.4569 | -106.4077
20.4606 | -106.4069
20.4643 | -106.4076
20.4679 | -106.4095
20.4716 | -106.411
20.4753 | -106.4123
20.4789 | -106.4134
20.4826 | -106.4132
20.4862 | -106.4133
20.4899 | -106.4136
20.4936 | -106.4148
20.4972 | -106.4163
20.5009 | -106.4181
20.5046 | -106.4201
20.5082 | -106.4233
20.5119 | -106.4258
20.5155 | -106.4274
20.5192 | -106.4287
20.5229 | -106.4309
20.5265 | -106.4326
20.5302 | -106.4334
20.5339 | -106.4348
20.5375 | -106.4376
20.5412 | -106.4403
20.5448 | -106.4421
20.5485 | -106.4461
20.5522 | -106.4506
20.5558 | -106.455
20.5595 | -106.4589
20.5632 | -106.4602
20.5668 | -106.4617
20.5705 | -106.4655
20.5741 | -106.4667
20.5778 | -106.469
20.5815 | -106.4721
20.5851 | -106.4733
20.5888 | -106.4719
20.5925 | -106.4719
20.5961 | -106.4698
20.5998 | -106.4684
20.6035 | -106.4678
20.6071 | -106.4667
20.6108 | -106.4679
20.6144 | -106.468
20.6181 | -106.4681
20.6218 | -106.4659
20.6254 | -106.4659
20.6291 | -106.4663
20.6328 | -106.4647
20.6364 | -106.4632
20.6401 | -106.4614
20.6437 | -106.459
20.6474 | -106.4564
20.6511 | -106.4557
20.6547 | -106.4547
20.6584 | -106.4534
20.6621 | -106.4499
20.6657 | -106.4466
20.6694 | -106.4427
20.673 | -106.4376
20.6767 | -106.4319
20.6804 | -106.4254
20.684 | -106.4203
20.6877 | -106.4165
20.6914 | -106.4109
20.695 | -106.4034
20.6987 | -106.3983
20.7023 | -106.3918
20.706 | -106.3848
20.7097 | -106.3785
20.7133 | -106.3716
20.717 | -106.362
20.7207 | -106.3533
20.7243 | -106.3452
20.728 | -106.3379
20.7317 | -106.3309
20.7353 | -106.3237
20.739 | -106.3174
20.7426 | -106.3119
20.7463 | -106.3092
20.75 | -106.3058
20.7536 | -106.3033
20.7573 | -106.2999
20.761 | -106.2975
20.7646 | -106.2961
20.7683 | -106.2955
20.7719 | -106.2946
20.7756 | -106.2952
20.7793 | -106.2961
20.7829 | -106.2984
20.7866 | -106.3011
20.7903 | -106.3035
20.7939 | -106.306
20.7976 | -106.3085
20.8012 | -106.3118
20.8049 | -106.3145
20.8086 | -106.3185
20.8122 | -106.3236
20.8159 | -106.3271
20.8196 | -106.3325
20.8232 | -106.3401
20.8269 | -106.3474
20.8305 | -106.3547
20.8342 | -106.3637
20.8379 | -106.3722
20.8415 | -106.3795
20.8452 | -106.3882
20.8489 | -106.398
20.8525 | -106.4081
20.8562 | -106.4193
20.8599 | -106.4306
20.8635 | -106.4417
20.8672 | -106.4541
20.8708 | -106.4653
20.8745 | -106.4763
20.8782 | -106.4877
20.8818 | -106.5005
20.8855 | -106.5119
20.8892 | -106.5223
20.8928 | -106.5347
20.8965 | -106.5466
20.9001 | -106.5564
20.9038 | -106.5645
20.9075 | -106.5719
20.9111 | -106.582
20.9148 | -106.5907
20.9185 | -106.6016
20.9221 | -106.6087
20.9258 | -106.6164
20.9294 | -106.621
20.9331 | -106.626
20.9368 | -106.6347
20.9404 | -106.6429
20.9441 | -106.6482
20.9478 | -106.6496
20.9514 | -106.6491
20.9551 | -106.6503
20.9587 | -106.6542
20.9624 | -106.6583
20.9661 | -106.6604
20.9697 | -106.6634
20.9734 | -106.6656
20.9771 | -106.6683
20.9807 | -106.6686
20.9844 | -106.6666
20.9881 | -106.6657
20.9917 | -106.6664
20.9954 | -106.6663
20.999 | -106.6659
21.0027 | -106.6674
21.0064 | -106.6696
21.01 | -106.6713
21.0137 | -106.6732
21.0174 | -106.6743
21.021 | -106.6726
21.0247 | -106.6719
21.0283 | -106.674
21.032 | -106.6755
21.0357 | -106.6754
21.0393 | -106.6749
21.043 | -106.6756
21.0467 | -106.6781
21.0503 | -106.6794
21.054 | -106.6816
21.0576 | -106.6851
21.0613 | -106.6873
21.065 | -106.6879
21.0686 | -106.6913
21.0723 | -106.694
21.076 | -106.6949
21.0796 | -106.6956
21.0833 | -106.6964
21.0869 | -106.6978
21.0906 | -106.7002
21.0943 | -106.7001
21.0979 | -106.6977
21.1016 | -106.696
21.1053 | -106.6968
21.1089 | -106.6977
21.1126 | -106.6992
21.1162 | -106.6986
21.1199 | -106.6956
21.1236 | -106.6923
21.1272 | -106.6874
21.1309 | -106.6803
21.1346 | -106.665
21.1382 | -106.656
21.1419 | -106.6477
21.1456 | -106.6364
21.1492 | -106.6252
21.1529 | -106.6133
21.1565 | -106.5998
21.1602 | -106.5868
21.1639 | -106.5718
21.1675 | -106.5542
21.1712 | -106.5391
21.1749 | -106.5238
21.1785 | -106.5072
21.1822 | -106.4882
21.1858 | -106.47
21.1895 | -106.4515
21.1932 | -106.4316
21.1968 | -106.4125
21.2005 | -106.3946
21.2042 | -106.3786
21.2078 | -106.3609
21.2115 | -106.3422
21.2151 | -106.3247
21.2188 | -106.3075
21.2225 | -106.2904
21.2261 | -106.2732
21.2298 | -106.2587
21.2335 | -106.2426
21.2371 | -106.2263
21.2408 | -106.2125
21.2444 | -106.2005
21.2481 | -106.1901
21.2518 | -106.1812
21.2554 | -106.172
21.2591 | -106.1626
21.2628 | -106.1529
21.2664 | -106.1448
21.2701 | -106.136
21.2738 | -106.1263
21.2774 | -106.1186
21.2811 | -106.1128
21.2847 | -106.109
21.2884 | -106.1038
21.2921 | -106.0985
21.2957 | -106.0931
21.2994 | -106.0885
21.3031 | -106.0828
21.3067 | -106.077
21.3104 | -106.0739
21.314 | -106.0723
21.3177 | -106.0695
21.3214 | -106.0696
21.325 | -106.0677
21.3287 | -106.0642
21.3324 | -106.0602
21.336 | -106.0573
21.3397 | -106.0568
21.3433 | -106.0533
21.347 | -106.05
21.3507 | -106.0484
21.3543 | -106.0456
21.358 | -106.0406
21.3617 | -106.0361
21.3653 | -106.0355
21.369 | -106.035
21.3726 | -106.0352
21.3763 | -106.0358
21.38 | -106.0352
21.3836 | -106.035
21.3873 | -106.0365
21.391 | -106.0349
21.3946 | -106.0352
21.3983 | -106.0384
21.4019 | -106.0404
21.4056 | -106.0391
21.4093 | -106.0368
21.4129 | -106.0354
21.4166 | -106.0346
21.4203 | -106.0336
21.4239 | -106.0334
21.4276 | -106.0331
21.4313 | -106.0322
21.4349 | -106.0333
21.4386 | -106.0345
21.4422 | -106.0353
21.4459 | -106.0353
21.4496 | -106.0356
21.4532 | -106.0347
21.4569 | -106.0352
21.4606 | -106.037
21.4642 | -106.0364
21.4679 | -106.0363
21.4715 | -106.0374
21.4752 | -106.0351
21.4789 | -106.0326
21.4825 | -106.0322
21.4862 | -106.0317
21.4899 | -106.0304
21.4935 | -106.0302
21.4972 | -106.0298
21.5008 | -106.0286
21.5045 | -106.0286
21.5082 | -106.0278
21.5118 | -106.0273
21.5155 | -106.0274
21.5192 | -106.0281
21.5228 | -106.028
21.5265 | -106.0282
21.5301 | -106.0285
21.5338 | -106.0308
21.5375 | -106.0332
21.5411 | -106.0342
21.5448 | -106.0344
21.5485 | -106.0357
21.5521 | -106.037
21.5558 | -106.0383
21.5595 | -106.0393
21.5631 | -106.0404
21.5668 | -106.0435
21.5704 | -106.0445
21.5741 | -106.0474
21.5778 | -106.0532
21.5814 | -106.0564
21.5851 | -106.0567
21.5888 | -106.0541
21.5924 | -106.0548
21.5961 | -106.0572
21.5997 | -106.06
21.6034 | -106.0637
21.6071 | -106.0688
21.6107 | -106.0728
21.6144 | -106.074
21.6181 | -106.0736
21.6217 | -106.0735
21.6254 | -106.0748
21.629 | -106.0775
21.6327 | -106.0809
21.6364 | -106.0846
21.64 | -106.0849
21.6437 | -106.0841
21.6474 | -106.0858
21.651 | -106.0873
21.6547 | -106.0871
21.6583 | -106.0877
21.662 | -106.0887
21.6657 | -106.0852
21.6693 | -106.0866
21.673 | -106.084
21.6767 | -106.0815
21.6803 | -106.0808
21.684 | -106.0812
21.6877 | -106.0818
21.6913 | -106.0808
21.695 | -106.0784
21.6986 | -106.0781
21.7023 | -106.0772
21.706 | -106.0758
21.7096 | -106.0746
21.7133 | -106.0753
21.717 | -106.0769
21.7206 | -106.0772
21.7243 | -106.0768
21.7279 | -106.0776
21.7316 | -106.0768
21.7353 | -106.0775
21.7389 | -106.08
21.7426 | -106.0797
21.7463 | -106.0807
21.7499 | -106.0841
21.7536 | -106.0854
21.7572 | -106.0857
21.7609 | -106.0867
21.7646 | -106.0869
21.7682 | -106.087
21.7719 | -106.0888
21.7756 | -106.0898
21.7792 | -106.0913
21.7829 | -106.0943
21.7865 | -106.0966
21.7902 | -106.0976
21.7939 | -106.0995
21.7975 | -106.102
21.8012 | -106.1033
21.8049 | -106.1005
21.8085 | -106.0997
21.8122 | -106.0991
21.8159 | -106.0979
21.8195 | -106.0958
21.8232 | -106.0934
21.8268 | -106.09
21.8305 | -106.0858
21.8342 | -106.0832
21.8378 | -106.079
21.8415 | -106.0733
21.8452 | -106.067
21.8488 | -106.0605
21.8525 | -106.0551
21.8561 | -106.0473
21.8598 | -106.0383
21.8635 | -106.0299
21.8671 | -106.0212
21.8708 | -106.0114
21.8745 | -106.0008
21.8781 | -105.9925
21.8818 | -105.981
21.8854 | -105.9691
21.8891 | -105.9585
21.8928 | -105.9501
21.8964 | -105.9437
21.9001 | -105.9335
21.9038 | -105.9205
21.9074 | -105.9079
21.9111 | -105.8959
21.9147 | -105.8857
21.9184 | -105.8754
21.9221 | -105.868
21.9257 | -105.8631
21.9294 | -105.8576
21.9331 | -105.8491
21.9367 | -105.8427
21.9404 | -105.838
21.9441 | -105.8334
21.9477 | -105.8314
21.9514 | -105.8306
21.955 | -105.8269
21.9587 | -105.8243
21.9624 | -105.8252
21.966 | -105.8266
21.9697 | -105.8291
21.9734 | -105.8328
21.977 | -105.8347
21.9807 | -105.8376
21.9843 | -105.8408
21.988 | -105.8419
21.9917 | -105.8439
21.9953 | -105.8485
21.999 | -105.8543
22.0027 | -105.8606
22.0063 | -105.8678
22.01 | -105.876
22.0136 | -105.8836
22.0173 | -105.8904
22.021 | -105.8957
22.0246 | -105.9003
22.0283 | -105.9048
22.032 | -105.9101
22.0356 | -105.9156
22.0393 | -105.9208
22.0429 | -105.927
22.0466 | -105.9332
22.0503 | -105.9387
22.0539 | -105.9436
22.0576 | -105.9457
22.0613 | -105.9469
22.0649 | -105.9493
22.0686 | -105.9525
22.0722 | -105.9538
22.0759 | -105.9555
22.0796 | -105.9576
22.0832 | -105.9588
22.0869 | -105.9583
22.0906 | -105.957
22.0942 | -105.956
22.0979 | -105.9551
22.1016 | -105.9546
22.1052 | -105.9524
22.1089 | -105.949
22.1125 | -105.9441
22.1162 | -105.9422
22.1199 | -105.9376
22.1235 | -105.9315
22.1272 | -105.927
22.1309 | -105.9236
22.1345 | -105.9195
22.1382 | -105.916
22.1418 | -105.9129
22.1455 | -105.9086
22.1492 | -105.9048
22.1528 | -105.9013
22.1565 | -105.8982
22.1602 | -105.8935
22.1638 | -105.8886
22.1675 | -105.8867
22.1711 | -105.8866
22.1748 | -105.8859
22.1785 | -105.8854
22.1821 | -105.8849
22.1858 | -105.8853
22.1895 | -105.8866
22.1931 | -105.8942
22.1968 | -105.8967
22.2004 | -105.8985
22.2041 | -105.9021
22.2078 | -105.9075
22.2114 | -105.9125
22.2151 | -105.9179
22.2188 | -105.925
22.2224 | -105.9323
22.2261 | -105.9392
22.2298 | -105.946
22.2334 | -105.9528
22.2371 | -105.959
22.2407 | -105.9657
22.2444 | -105.9752
22.2481 | -105.9857
22.2517 | -105.9943
22.2554 | -106.0034
22.2591 | -106.0139
22.2627 | -106.0239
22.2664 | -106.0342
22.27 | -106.046
22.2737 | -106.0561
22.2774 | -106.0639
22.281 | -106.073
22.2847 | -106.0821
22.2884 | -106.0919
22.292 | -106.1011
22.2957 | -106.1085
22.2993 | -106.1172
22.303 | -106.1278
22.3067 | -106.1353
22.3103 | -106.1395
22.314 | -106.1437
22.3177 | -106.1486
22.3213 | -106.1512
22.325 | -106.1522
22.3286 | -106.1526
22.3323 | -106.1531
22.336 | -106.1538
22.3396 | -106.1545
22.3433 | -106.1549
22.347 | -106.1554
22.3506 | -106.1563
22.3543 | -106.1556
22.358 | -106.1517
22.3616 | -106.148
22.3653 | -106.1417
22.3689 | -106.1357
22.3726 | -106.1307
22.3763 | -106.1257
22.3799 | -106.12
22.3836 | -106.1133
22.3873 | -106.1062
22.3909 | -106.0985
22.3946 | -106.0883
22.3982 | -106.078
22.4019 | -106.0708
22.4056 | -106.0625
22.4092 | -106.0508
22.4129 | -106.0425
22.4166 | -106.0319
22.4202 | -106.0227
22.4239 | -106.0156
22.4275 | -106.0088
22.4312 | -106.0006
22.4349 | -105.9926
22.4385 | -105.9837
22.4422 | -105.9731
22.4459 | -105.9633
22.4495 | -105.9533
22.4532 | -105.9445
22.4568 | -105.9354
22.4605 | -105.9276
22.4642 | -105.9179
22.4678 | -105.9072
22.4715 | -105.8993
22.4752 | -105.8938
22.4788 | -105.8875
22.4825 | -105.879
22.4861 | -105.8709
22.4898 | -105.8615
22.4935 | -105.8527
22.4971 | -105.8469
22.5008 | -105.8393
22.5045 | -105.8309
22.5081 | -105.8253
22.5118 | -105.8194
22.5155 | -105.8146
22.5191 | -105.81
22.5228 | -105.8054
22.5264 | -105.7985
22.5301 | -105.7931
22.5338 | -105.7881
22.5374 | -105.7821
22.5411 | -105.7765
22.5448 | -105.7712
22.5484 | -105.7652
22.5521 | -105.7611
22.5557 | -105.7566
22.5594 | -105.7489
22.5631 | -105.7409
22.5667 | -105.733
22.5704 | -105.7253
22.5741 | -105.7176
22.5777 | -105.7096
22.5814 | -105.7013
22.585 | -105.6914
22.5887 | -105.6829
22.5924 | -105.6749
22.596 | -105.6672
22.5997 | -105.6595
22.6034 | -105.6508
22.607 | -105.6417
22.6107 | -105.6338
22.6143 | -105.6267
22.618 | -105.6194
22.6217 | -105.61
22.6253 | -105.601
22.629 | -105.5942
22.6327 | -105.5885
22.6363 | -105.5822
22.64 | -105.5749
22.6437 | -105.5699
22.6473 | -105.5664
22.651 | -105.5623
22.6546 | -105.5583
22.6583 | -105.5527
22.662 | -105.5458
22.6656 | -105.5413
22.6693 | -105.5385
22.673 | -105.536
22.6766 | -105.5334
22.6803 | -105.5305
22.6839 | -105.5299
22.6876 | -105.5321
22.6913 | -105.5358
22.6949 | -105.5383
22.6986 | -105.5421
22.7023 | -105.5463
22.7059 | -105.5514
22.7096 | -105.5576
22.7132 | -105.5652
22.7169 | -105.5728
22.7206 | -105.5857
22.7242 | -105.5927
22.7279 | -105.6015
22.7316 | -105.6132
22.7352 | -105.6226
22.7389 | -105.6347
22.7425 | -105.6467
22.7462 | -105.6562
22.7499 | -105.6675
22.7535 | -105.6799
22.7572 | -105.6918
22.7609 | -105.7045
22.7645 | -105.7191
22.7682 | -105.7329
22.7719 | -105.7472
22.7755 | -105.7605
22.7792 | -105.7758
22.7828 | -105.7906
22.7865 | -105.8068
22.7902 | -105.8228
22.7938 | -105.8369
22.7975 | -105.8489
22.8012 | -105.8611
22.8048 | -105.8754
22.8085 | -105.8893
22.8121 | -105.9041
22.8158 | -105.9184
22.8195 | -105.9302
22.8231 | -105.9424
22.8268 | -105.9548
22.8305 | -105.9646
22.8341 | -105.9755
22.8378 | -105.9861
22.8414 | -105.9943
22.8451 | -106.0019
22.8488 | -106.0109
22.8524 | -106.0181
22.8561 | -106.0242
22.8598 | -106.0318
22.8634 | -106.037
22.8671 | -106.0424
22.8707 | -106.0456
22.8744 | -106.0471
22.8781 | -106.0491
22.8817 | -106.0534
22.8854 | -106.0544
22.8891 | -106.0554
22.8927 | -106.0559
22.8964 | -106.0558
22.9001 | -106.0542
22.9037 | -106.0535
22.9074 | -106.0523
22.911 | -106.0485
22.9147 | -106.0447
22.9184 | -106.0389
22.922 | -106.0327
22.9257 | -106.028
22.9294 | -106.0219
22.933 | -106.0147
22.9367 | -106.0079
22.9403 | -105.9992
22.944 | -105.992
22.9477 | -105.9857
22.9513 | -105.9774
22.955 | -105.9678
22.9587 | -105.9578
22.9623 | -105.9505
22.966 | -105.9412
22.9696 | -105.9318
22.9733 | -105.9223
22.977 | -105.9137
22.9806 | -105.9063
22.9843 | -105.9003
22.988 | -105.8938
22.9916 | -105.8873
22.9953 | -105.8819
22.9989 | -105.8766
23.0026 | -105.8731
23.0063 | -105.8697
23.0099 | -105.8645
23.0136 | -105.8593
23.0173 | -105.857
23.0209 | -105.8546
23.0246 | -105.8529
23.0282 | -105.8536
23.0319 | -105.8537
23.0356 | -105.8552
23.0392 | -105.8591
23.0429 | -105.8633
23.0466 | -105.8675
23.0502 | -105.8733
23.0539 | -105.8776
23.0576 | -105.8819
23.0612 | -105.8885
23.0649 | -105.8959
23.0685 | -105.9025
23.0722 | -105.9116
23.0759 | -105.9203
23.0795 | -105.9298
23.0832 | -105.9396
23.0869 | -105.9492
23.0905 | -105.9576
23.0942 | -105.9663
23.0978 | -105.9743
23.1015 | -105.9816
23.1052 | -105.9885
23.1088 | -105.9947
23.1125 | -106.0008
23.1162 | -106.0081
23.1198 | -106.0151
23.1235 | -106.0224
23.1271 | -106.0298
23.1308 | -106.0353
23.1345 | -106.0401
23.1381 | -106.0425
23.1418 | -106.0464
23.1455 | -106.051
23.1491 | -106.0565
23.1528 | -106.0605
23.1564 | -106.0626
23.1601 | -106.0637
23.1638 | -106.0659
23.1674 | -106.0685
23.1711 | -106.0694
23.1748 | -106.0694
23.1784 | -106.0693
23.1821 | -106.072
23.1858 | -106.073
23.1894 | -106.0748
23.1931 | -106.0763
23.1967 | -106.077
23.2004 | -106.0774
23.2041 | -106.078
23.2077 | -106.078
23.2114 | -106.0777
23.2151 | -106.0771
23.2187 | -106.0763
23.2224 | -106.0755
23.226 | -106.0757
23.2297 | -106.0758
23.2334 | -106.0757
23.237 | -106.0759
23.2407 | -106.0767
23.2444 | -106.08
23.248 | -106.0854
23.2517 | -106.0887
23.2553 | -106.0918
23.259 | -106.0944
23.2627 | -106.0988
23.2663 | -106.1041
23.27 | -106.1095
23.2737 | -106.1135
23.2773 | -106.1179
23.281 | -106.1234
23.2846 | -106.1287
23.2883 | -106.1358
23.292 | -106.141
23.2956 | -106.147
23.2993 | -106.1532
23.303 | -106.1594
23.3066 | -106.1663
23.3103 | -106.1738
23.314 | -106.1803
23.3176 | -106.1854
23.3213 | -106.1907
23.3249 | -106.1967
23.3286 | -106.2039
23.3323 | -106.2092
23.3359 | -106.2143
23.3396 | -106.2206
23.3433 | -106.2268
23.3469 | -106.2315
23.3506 | -106.2367
23.3542 | -106.2403
23.3579 | -106.2424
23.3616 | -106.2466
23.3652 | -106.2507
23.3689 | -106.2532
23.3726 | -106.255
23.3762 | -106.2561
23.3799 | -106.2577
23.3835 | -106.2619
23.3872 | -106.2642
23.3909 | -106.2645
23.3945 | -106.2648
23.3982 | -106.2677
23.4019 | -106.2698
23.4055 | -106.2699
23.4092 | -106.2701
23.4128 | -106.2711
23.4165 | -106.2712
23.4202 | -106.2704
23.4238 | -106.2688
23.4275 | -106.2687
23.4312 | -106.2702
23.4348 | -106.2701
23.4385 | -106.271
23.4422 | -106.2717
23.4458 | -106.2709
23.4495 | -106.2686
23.4531 | -106.2656
23.4568 | -106.2622
23.4605 | -106.26
23.4641 | -106.258
23.4678 | -106.2589
23.4715 | -106.2573
23.4751 | -106.2554
23.4788 | -106.2528
23.4824 | -106.2509
23.4861 | -106.2505
23.4898 | -106.2506
23.4934 | -106.2494
23.4971 | -106.2489
23.5008 | -106.2467
23.5044 | -106.2447
23.5081 | -106.244
23.5117 | -106.2423
23.5154 | -106.2402
23.5191 | -106.2378
23.5227 | -106.2346
23.5264 | -106.2331
23.5301 | -106.2316
23.5337 | -106.2292
23.5374 | -106.2288
23.541 | -106.2297
23.5447 | -106.2291
23.5484 | -106.2277
23.552 | -106.2272
23.5557 | -106.2268
23.5594 | -106.2279
23.563 | -106.2299
23.5667 | -106.2314
23.5703 | -106.2316
23.574 | -106.2317
23.5777 | -106.2331
23.5813 | -106.2346
23.585 | -106.2357
23.5887 | -106.2374
23.5923 | -106.2423
23.596 | -106.2458
23.5997 | -106.2493
23.6033 | -106.2546
23.607 | -106.2593
23.6106 | -106.2617
23.6143 | -106.2643
23.618 | -106.2699
23.6216 | -106.2765
23.6253 | -106.2839
23.629 | -106.2904
23.6326 | -106.2974
23.6363 | -106.3045
23.6399 | -106.3095
23.6436 | -106.3138
23.6473 | -106.3207
23.6509 | -106.3282
23.6546 | -106.3351
23.6583 | -106.3405
23.6619 | -106.3483
23.6656 | -106.3578
23.6692 | -106.3662
23.6729 | -106.3733
23.6766 | -106.381
23.6802 | -106.3885
23.6839 | -106.3952
23.6876 | -106.4037
23.6912 | -106.4131
23.6949 | -106.4211
23.6985 | -106.4289
23.7022 | -106.4346
23.7059 | -106.4417
23.7095 | -106.4509
23.7132 | -106.4586
23.7169 | -106.463
23.7205 | -106.4681
23.7242 | -106.4731
23.7279 | -106.4787
23.7315 | -106.4821
23.7352 | -106.4838
23.7388 | -106.4883
23.7425 | -106.4916
23.7462 | -106.4956
23.7498 | -106.4986
23.7535 | -106.5007
23.7572 | -106.501
23.7608 | -106.5007
23.7645 | -106.4992
23.7681 | -106.4975
23.7718 | -106.496
23.7755 | -106.4946
23.7791 | -106.4923
23.7828 | -106.4908
23.7865 | -106.4902
23.7901 | -106.4906
23.7938 | -106.4912
23.7974 | -106.4893
23.8011 | -106.4867
23.8048 | -106.4833
23.8084 | -106.4799
23.8121 | -106.4774
23.8158 | -106.4754
23.8194 | -106.4714
23.8231 | -106.4674
23.8267 | -106.4642
23.8304 | -106.4606
23.8341 | -106.4568
23.8377 | -106.4553
23.8414 | -106.4539
23.8451 | -106.4528
23.8487 | -106.451
23.8524 | -106.4492
23.8561 | -106.4482
23.8597 | -106.4481
23.8634 | -106.4476
23.867 | -106.448
23.8707 | -106.4479
23.8744 | -106.4468
23.878 | -106.4456
23.8817 | -106.4454
23.8854 | -106.4474
23.889 | -106.4483
23.8927 | -106.4502
23.8963 | -106.4515
23.9 | -106.4518
23.9037 | -106.4542
23.9073 | -106.4588
23.911 | -106.4609
23.9147 | -106.4632
23.9183 | -106.467
23.922 | -106.4715
23.9256 | -106.4754
23.9293 | -106.4791
23.933 | -106.4843
23.9366 | -106.4885
23.9403 | -106.4919
23.944 | -106.4963
23.9476 | -106.5025
23.9513 | -106.5076
23.9549 | -106.5117
23.9586 | -106.5167
23.9623 | -106.5229
23.9659 | -106.529
23.9696 | -106.5325
23.9733 | -106.5379
23.9769 | -106.5448
23.9806 | -106.5518
23.9842 | -106.5571
23.9879 | -106.5623
23.9916 | -106.5677
23.9952 | -106.5749
23.9989 | -106.5806
24.0026 | -106.5857
24.0062 | -106.591
24.0099 | -106.596
24.0136 | -106.6034
24.0172 | -106.6114
24.0209 | -106.6203
24.0245 | -106.6296
24.0282 | -106.6373
24.0319 | -106.6448
24.0355 | -106.6523
24.0392 | -106.6617
24.0429 | -106.6711
24.0465 | -106.6788
24.0502 | -106.687
24.0538 | -106.698
24.0575 | -106.7084
24.0612 | -106.7191
24.0648 | -106.7292
24.0685 | -106.7411
24.0722 | -106.755
24.0758 | -106.7675
24.0795 | -106.7803
24.0831 | -106.7898
24.0868 | -106.8013
24.0905 | -106.8136
24.0941 | -106.8282
24.0978 | -106.8399
24.1015 | -106.8503
24.1051 | -106.8603
24.1088 | -106.8692
24.1124 | -106.8779
24.1161 | -106.887
24.1198 | -106.8955
24.1234 | -106.9051
24.1271 | -106.9146
24.1308 | -106.9229
24.1344 | -106.9279
24.1381 | -106.9347
24.1418 | -106.9414
24.1454 | -106.9474
24.1491 | -106.9527
24.1527 | -106.9552
24.1564 | -106.9586
24.1601 | -106.9648
24.1637 | -106.9693
24.1674 | -106.9716
24.1711 | -106.973
24.1747 | -106.9751
24.1784 | -106.9763
24.182 | -106.9758
24.1857 | -106.9743
24.1894 | -106.9736
24.193 | -106.9722
24.1967 | -106.9707
24.2004 | -106.9684
24.204 | -106.9644
24.2077 | -106.9596
24.2113 | -106.957
24.215 | -106.9531
24.2187 | -106.95
24.2223 | -106.9464
24.226 | -106.9425
24.2297 | -106.9393
24.2333 | -106.9357
24.237 | -106.9315
24.2406 | -106.9291
24.2443 | -106.9262
24.248 | -106.921
24.2516 | -106.9181
24.2553 | -106.9151
24.259 | -106.9129
24.2626 | -106.9104
24.2663 | -106.9079
24.27 | -106.904
24.2736 | -106.8993
24.2773 | -106.896
24.2809 | -106.8951
24.2846 | -106.8946
24.2883 | -106.8942
24.2919 | -106.8933
24.2956 | -106.8895
24.2993 | -106.8863
24.3029 | -106.8845
24.3066 | -106.8838
24.3102 | -106.8822
24.3139 | -106.8821
24.3176 | -106.8801
24.3212 | -106.8779
24.3249 | -106.875
24.3286 | -106.8718
24.3322 | -106.8693
24.3359 | -106.8658
24.3395 | -106.8617
24.3432 | -106.859
24.3469 | -106.855
24.3505 | -106.85
24.3542 | -106.8463
24.3579 | -106.8399
24.3615 | -106.8332
24.3652 | -106.8252
24.3688 | -106.8172
24.3725 | -106.81
24.3762 | -106.8032
24.3798 | -106.7948
24.3835 | -106.7885
24.3872 | -106.7811
24.3908 | -106.7715
24.3945 | -106.762
24.3982 | -106.7525
24.4018 | -106.7437
24.4055 | -106.7327
24.4091 | -106.7229
24.4128 | -106.713
24.4165 | -106.7024
24.4201 | -106.6921
24.4238 | -106.6833
24.4275 | -106.6748
24.4311 | -106.6648
24.4348 | -106.6547
24.4384 | -106.6452
24.4421 | -106.636
24.4458 | -106.6285
24.4494 | -106.6219
24.4531 | -106.616
24.4568 | -106.6094
24.4604 | -106.6037
24.4641 | -106.5981
24.4677 | -106.5921
24.4714 | -106.5864
24.4751 | -106.5806
24.4787 | -106.5756
24.4824 | -106.5707
24.4861 | -106.5677
24.4897 | -106.5654
24.4934 | -106.5639
24.497 | -106.5628
24.5007 | -106.5612
24.5044 | -106.5598
24.508 | -106.5594
24.5117 | -106.5567
24.5154 | -106.5547
24.519 | -106.5541
24.5227 | -106.5543
24.5264 | -106.5542
24.53 | -106.5548
24.5337 | -106.5547
24.5373 | -106.5549
24.541 | -106.5563
24.5447 | -106.5587
24.5483 | -106.5616
24.552 | -106.5653
24.5557 | -106.567
24.5593 | -106.5682
24.563 | -106.5683
24.5666 | -106.5691
24.5703 | -106.5685
24.574 | -106.5641
24.5776 | -106.5597
24.5813 | -106.5569
24.585 | -106.553
24.5886 | -106.5417
24.5923 | -106.529
24.5959 | -106.5151
24.5996 | -106.5036
24.6033 | -106.491
24.6069 | -106.4765
24.6106 | -106.459
24.6143 | -106.4436
24.6179 | -106.424
24.6216 | -106.4024
24.6252 | -106.3801
24.6289 | -106.3578
24.6326 | -106.3338
24.6362 | -106.3092
24.6399 | -106.2845
24.6436 | -106.2569
24.6472 | -106.2274
24.6509 | -106.1974
24.6545 | -106.1663
24.6582 | -106.1336
24.6619 | -106.0997
24.6655 | -106.0661
24.6692 | -106.0304
24.6729 | -105.9928
24.6765 | -105.9548
24.6802 | -105.9158
24.6839 | -105.8772
24.6875 | -105.839
24.6912 | -105.7987
24.6948 | -105.7597
24.6985 | -105.7219
24.7022 | -105.6834
24.7058 | -105.6448
24.7095 | -105.6062
24.7132 | -105.5699
24.7168 | -105.5336
24.7205 | -105.4981
24.7241 | -105.4645
24.7278 | -105.4336
24.7315 | -105.403
24.7351 | -105.3741
24.7388 | -105.3477
24.7425 | -105.323
24.7461 | -105.2986
24.7498 | -105.275
24.7534 | -105.2554
24.7571 | -105.2369
24.7608 | -105.221
24.7644 | -105.2082
24.7681 | -105.1976
24.7718 | -105.1887
24.7754 | -105.1821
24.7791 | -105.1759
24.7827 | -105.175
24.7864 | -105.1759
24.7901 | -105.1791
24.7937 | -105.1854
24.7974 | -105.1931
24.8011 | -105.2024
24.8047 | -105.2147
24.8084 | -105.2277
24.8121 | -105.2419
24.8157 | -105.2594
24.8194 | -105.2799
24.823 | -105.3019
24.8267 | -105.3265
24.8304 | -105.3522
24.834 | -105.3821
24.8377 | -105.4065
24.8414 | -105.4313
24.845 | -105.458
24.8487 | -105.4857
24.8523 | -105.5127
24.856 | -105.5393
24.8597 | -105.5672
24.8633 | -105.598
24.867 | -105.6286
24.8707 | -105.656
24.8743 | -105.6814
24.878 | -105.7102
24.8816 | -105.7371
24.8853 | -105.7631
24.889 | -105.7889
24.8926 | -105.8138
24.8963 | -105.838
24.9 | -105.8615
24.9036 | -105.8831
24.9073 | -105.9047
24.9109 | -105.9271
24.9146 | -105.9487
24.9183 | -105.9693
24.9219 | -105.9896
24.9256 | -106.0116
24.9293 | -106.0313
24.9329 | -106.0487
24.9366 | -106.0649
24.9402 | -106.0853
24.9439 | -106.1047
24.9476 | -106.1221
24.9512 | -106.135
24.9549 | -106.1458
24.9586 | -106.1574
24.9622 | -106.1703
24.9659 | -106.1797
24.9696 | -106.1909
24.9732 | -106.1998
24.9769 | -106.2073
24.9805 | -106.2142
24.9842 | -106.2197
24.9879 | -106.2242
24.9915 | -106.2274
24.9952 | -106.2314
24.9989 | -106.236
25.0025 | -106.2375
25.0062 | -106.2403
25.0098 | -106.2429
25.0135 | -106.2456
25.0172 | -106.2479
25.0208 | -106.2514
25.0245 | -106.2548
25.0282 | -106.258
25.0318 | -106.2604
25.0355 | -106.2626
25.0391 | -106.2655
25.0428 | -106.2675
25.0465 | -106.269
25.0501 | -106.2708
25.0538 | -106.273
25.0575 | -106.2758
25.0611 | -106.2773
25.0648 | -106.278
25.0684 | -106.2795
25.0721 | -106.2806
25.0758 | -106.2805
25.0794 | -106.2796
25.0831 | -106.2783
25.0868 | -106.2789
25.0904 | -106.2795
25.0941 | -106.2791
25.0978 | -106.2795
25.1014 | -106.281
25.1051 | -106.2808
25.1087 | -106.2789
25.1124 | -106.2784
25.1161 | -106.2773
25.1197 | -106.2764
25.1234 | -106.2758
25.1271 | -106.276
25.1307 | -106.2782
25.1344 | -106.2764
25.138 | -106.2737
25.1417 | -106.2723
25.1454 | -106.2692
25.149 | -106.2665
25.1527 | -106.2638
25.1564 | -106.2616
25.16 | -106.2599
25.1637 | -106.2579
25.1673 | -106.2553
25.171 | -106.2517
25.1747 | -106.2494
25.1783 | -106.2466
25.182 | -106.242
25.1857 | -106.2395
25.1893 | -106.2387
25.193 | -106.2356
25.1966 | -106.2327
25.2003 | -106.2288
25.204 | -106.2261
25.2076 | -106.2252
25.2113 | -106.2246
25.215 | -106.223
25.2186 | -106.2217
25.2223 | -106.2202
25.226 | -106.2185
25.2296 | -106.2192
25.2333 | -106.2174
25.2369 | -106.2174
25.2406 | -106.2173
25.2443 | -106.216
25.2479 | -106.2151
25.2516 | -106.2158
25.2553 | -106.2165
25.2589 | -106.2157
25.2626 | -106.2141
25.2662 | -106.2126
25.2699 | -106.2126
25.2736 | -106.2113
25.2772 | -106.2104
25.2809 | -106.2108
25.2846 | -106.2131
25.2882 | -106.2136
25.2919 | -106.2137
25.2955 | -106.2144
25.2992 | -106.2151
25.3029 | -106.2152
25.3065 | -106.2157
25.3102 | -106.2168
25.3139 | -106.2161
25.3175 | -106.217
25.3212 | -106.2177
25.3248 | -106.2166
25.3285 | -106.2171
25.3322 | -106.2181
25.3358 | -106.2181
25.3395 | -106.2185
25.3432 | -106.2207
25.3468 | -106.2221
25.3505 | -106.2207
25.3542 | -106.2182
25.3578 | -106.2174
25.3615 | -106.2142
25.3651 | -106.2127
25.3688 | -106.2117
25.3725 | -106.2109
25.3761 | -106.2093
25.3798 | -106.2064
25.3835 | -106.2048
25.3871 | -106.2016
25.3908 | -106.1992
25.3944 | -106.1971
25.3981 | -106.1932
25.4018 | -106.1894
25.4054 | -106.1854
25.4091 | -106.1816
25.4128 | -106.1778
25.4164 | -106.1729
25.4201 | -106.1696
25.4237 | -106.1662
25.4274 | -106.1622
25.4311 | -106.1584
25.4347 | -106.1554
25.4384 | -106.1526
25.4421 | -106.1487
25.4457 | -106.1442
25.4494 | -106.1412
25.453 | -106.1379
25.4567 | -106.1338
25.4604 | -106.1301
25.464 | -106.127
25.4677 | -106.1243
25.4714 | -106.1211
25.475 | -106.119
25.4787 | -106.1163
25.4824 | -106.1127
25.486 | -106.1099
25.4897 | -106.1068
25.4933 | -106.1036
25.497 | -106.1023
25.5007 | -106.1002
25.5043 | -106.0994
25.508 | -106.0985
25.5117 | -106.0966
25.5153 | -106.0949
25.519 | -106.0943
25.5226 | -106.0947
25.5263 | -106.094
25.53 | -106.0932
25.5336 | -106.0927
25.5373 | -106.0927
25.541 | -106.0951
25.5446 | -106.098
25.5483 | -106.0998
25.5519 | -106.1015
25.5556 | -106.103
25.5593 | -106.1039
25.5629 | -106.1036
25.5666 | -106.1034
25.5703 | -106.1046
25.5739 | -106.1048
25.5776 | -106.1039
25.5812 | -106.1041
25.5849 | -106.1056
25.5886 | -106.1051
25.5922 | -106.104
25.5959 | -106.1031
25.5996 | -106.1028
25.6032 | -106.1035
25.6069 | -106.1038
25.6105 | -106.1034
25.6142 | -106.1034
25.6179 | -106.1031
25.6215 | -106.1029
25.6252 | -106.1032
25.6289 | -106.1023
25.6325 | -106.1015
25.6362 | -106.1008
25.6399 | -106.0976
25.6435 | -106.0938
25.6472 | -106.0923
25.6508 | -106.0907
25.6545 | -106.0884
25.6582 | -106.0864
25.6618 | -106.0835
25.6655 | -106.0785
25.6692 | -106.075
25.6728 | -106.0729
25.6765 | -106.0692
25.6801 | -106.0655
25.6838 | -106.0639
25.6875 | -106.0611
25.6911 | -106.0568
25.6948 | -106.0542
25.6985 | -106.0513
25.7021 | -106.0467
25.7058 | -106.0434
25.7094 | -106.0412
25.7131 | -106.0397
25.7168 | -106.0399
25.7204 | -106.0378
25.7241 | -106.0367
25.7278 | -106.0363
25.7314 | -106.0352
25.7351 | -106.0334
25.7387 | -106.0337
25.7424 | -106.0352
25.7461 | -106.034
25.7497 | -106.0321
25.7534 | -106.0309
25.7571 | -106.0315
25.7607 | -106.0328
25.7644 | -106.0356
25.7681 | -106.0387
25.7717 | -106.0418
25.7754 | -106.045
25.779 | -106.0486
25.7827 | -106.0532
25.7864 | -106.058
25.79 | -106.0637
25.7937 | -106.0694
25.7974 | -106.0747
25.801 | -106.0811
25.8047 | -106.0886
25.8083 | -106.0957
25.812 | -106.1016
25.8157 | -106.1082
25.8193 | -106.1171
25.823 | -106.1246
25.8267 | -106.132
25.8303 | -106.1416
25.834 | -106.1493
25.8376 | -106.1573
25.8413 | -106.1662
25.845 | -106.1769
25.8486 | -106.1875
25.8523 | -106.1993
25.856 | -106.2105
25.8596 | -106.2194
25.8633 | -106.2281
25.8669 | -106.2367
25.8706 | -106.2474
25.8743 | -106.2593
25.8779 | -106.2702
25.8816 | -106.2814
25.8853 | -106.2922
25.8889 | -106.3047
25.8926 | -106.3163
25.8962 | -106.3273
25.8999 | -106.3371
25.9036 | -106.3464
25.9072 | -106.3548
25.9109 | -106.3649
25.9146 | -106.3745
25.9182 | -106.3848
25.9219 | -106.3955
25.9256 | -106.4064
25.9292 | -106.4174
25.9329 | -106.427
25.9365 | -106.4364
25.9402 | -106.4472
25.9439 | -106.4578
25.9475 | -106.4651
25.9512 | -106.472
25.9549 | -106.4801
25.9585 | -106.488
25.9622 | -106.4957
25.9658 | -106.5035
25.9695 | -106.5101
25.9732 | -106.5143
25.9768 | -106.5173
25.9805 | -106.521
25.9842 | -106.5259
25.9878 | -106.5294
25.9915 | -106.5306
25.9951 | -106.5317
25.9988 | -106.5338
26.0025 | -106.5348
26.0061 | -106.5365
26.0098 | -106.5373
26.0135 | -106.535
26.0171 | -106.5329
26.0208 | -106.5316
26.0244 | -106.5288
26.0281 | -106.5257
26.0318 | -106.5236
26.0354 | -106.5193
26.0391 | -106.5119
26.0428 | -106.5048
26.0464 | -106.498
26.0501 | -106.4916
26.0538 | -106.4862
26.0574 | -106.4805
26.0611 | -106.4735
26.0647 | -106.466
26.0684 | -106.4591
26.0721 | -106.4528
26.0757 | -106.4443
26.0794 | -106.436
26.0831 | -106.4256
26.0867 | -106.4159
26.0904 | -106.4053
26.094 | -106.3944
26.0977 | -106.384
26.1014 | -106.3746
26.105 | -106.3642
26.1087 | -106.3546
26.1124 | -106.3482
26.116 | -106.342
26.1197 | -106.3344
26.1233 | -106.3259
26.127 | -106.3171
26.1307 | -106.3087
26.1343 | -106.2993
26.138 | -106.2894
26.1417 | -106.2817
26.1453 | -106.2751
26.149 | -106.2671
26.1526 | -106.2608
26.1563 | -106.2527
26.16 | -106.2458
26.1636 | -106.2404
26.1673 | -106.2354
26.171 | -106.2301
26.1746 | -106.2252
26.1783 | -106.2214
26.182 | -106.219
26.1856 | -106.2152
26.1893 | -106.2133
26.1929 | -106.2128
26.1966 | -106.2114
26.2003 | -106.211
26.2039 | -106.2104
26.2076 | -106.2099
26.2113 | -106.209
26.2149 | -106.2096
26.2186 | -106.2107
26.2222 | -106.2131
26.2259 | -106.2158
26.2296 | -106.219
26.2332 | -106.2227
26.2369 | -106.2262
26.2406 | -106.2303
26.2442 | -106.2345
26.2479 | -106.2377
26.2515 | -106.242
26.2552 | -106.2466
26.2589 | -106.2514
26.2625 | -106.258
26.2662 | -106.2652
26.2699 | -106.2728
26.2735 | -106.2797
26.2772 | -106.2859
26.2808 | -106.2923
26.2845 | -106.2997
26.2882 | -106.3067
26.2918 | -106.3145
26.2955 | -106.323
26.2992 | -106.3323
26.3028 | -106.3412
26.3065 | -106.3505
26.3102 | -106.359
26.3138 | -106.3663
26.3175 | -106.376
26.3211 | -106.3863
26.3248 | -106.3946
26.3285 | -106.4028
26.3321 | -106.4097
26.3358 | -106.4161
26.3395 | -106.4232
26.3431 | -106.4313
26.3468 | -106.4383
26.3504 | -106.445
26.3541 | -106.451
26.3578 | -106.4575
26.3614 | -106.4637
26.3651 | -106.4695
26.3688 | -106.4733
26.3724 | -106.4769
26.3761 | -106.4807
26.3797 | -106.4835
26.3834 | -106.4874
26.3871 | -106.4921
26.3907 | -106.4944
26.3944 | -106.4955
26.3981 | -106.4987
26.4017 | -106.5009
26.4054 | -106.5026
26.409 | -106.5041
26.4127 | -106.5038
26.4164 | -106.504
26.42 | -106.5027
26.4237 | -106.5022
26.4274 | -106.5009
26.431 | -106.499
26.4347 | -106.4971
26.4383 | -106.4949
26.442 | -106.4928
26.4457 | -106.4907
26.4493 | -106.4903
26.453 | -106.4897
26.4567 | -106.4899
26.4603 | -106.4904
26.464 | -106.4896
26.4677 | -106.4883
26.4713 | -106.4868
26.475 | -106.485
26.4786 | -106.4823
26.4823 | -106.4798
26.486 | -106.4774
26.4896 | -106.4752
26.4933 | -106.4728
26.497 | -106.4698
26.5006 | -106.4679
26.5043 | -106.4662
26.5079 | -106.4638
26.5116 | -106.463
26.5153 | -106.461
26.5189 | -106.4587
26.5226 | -106.4577
26.5263 | -106.4539
26.5299 | -106.4531
26.5336 | -106.452
26.5372 | -106.4482
26.5409 | -106.4451
26.5446 | -106.4438
26.5482 | -106.4417
26.5519 | -106.4412
26.5556 | -106.4398
26.5592 | -106.44
26.5629 | -106.4404
26.5665 | -106.4391
26.5702 | -106.437
26.5739 | -106.436
26.5775 | -106.4363
26.5812 | -106.4375
26.5849 | -106.4382
26.5885 | -106.4384
26.5922 | -106.4402
26.5959 | -106.4421
26.5995 | -106.4434
26.6032 | -106.4458
26.6068 | -106.4469
26.6105 | -106.4483
26.6142 | -106.4498
26.6178 | -106.4513
26.6215 | -106.4542
26.6252 | -106.4573
26.6288 | -106.4595
26.6325 | -106.4623
26.6361 | -106.4639
26.6398 | -106.4667
26.6435 | -106.4683
26.6471 | -106.4709
26.6508 | -106.4748
26.6545 | -106.4754
26.6581 | -106.4751
26.6618 | -106.4759
26.6654 | -106.4767
26.6691 | -106.4764
26.6728 | -106.4761
26.6764 | -106.4766
26.6801 | -106.4765
26.6838 | -106.477
26.6874 | -106.4787
26.6911 | -106.479
26.6947 | -106.4789
26.6984 | -106.48
26.7021 | -106.479
26.7057 | -106.4785
26.7094 | -106.4768
26.7131 | -106.4743
26.7167 | -106.4733
26.7204 | -106.4727
26.7241 | -106.4715
26.7277 | -106.4699
26.7314 | -106.4666
26.735 | -106.4647
26.7387 | -106.4625
26.7424 | -106.46
26.746 | -106.4575
26.7497 | -106.4546
26.7534 | -106.454
26.757 | -106.451
26.7607 | -106.4458
26.7643 | -106.4419
26.768 | -106.4386
26.7717 | -106.4336
26.7753 | -106.4273
26.779 | -106.4221
26.7827 | -106.4196
26.7863 | -106.4169
26.79 | -106.4133
26.7936 | -106.4075
26.7973 | -106.4018
26.801 | -106.3967
26.8046 | -106.3904
26.8083 | -106.3844
26.812 | -106.3809
26.8156 | -106.3771
26.8193 | -106.372
26.8229 | -106.3668
26.8266 | -106.3611
26.8303 | -106.3538
26.8339 | -106.3453
26.8376 | -106.3374
26.8413 | -106.3318
26.8449 | -106.3273
26.8486 | -106.3229
26.8523 | -106.3184
26.8559 | -106.3129
26.8596 | -106.3067
26.8632 | -106.3003
26.8669 | -106.294
26.8706 | -106.2879
26.8742 | -106.2811
26.8779 | -106.2743
26.8816 | -106.2699
26.8852 | -106.2679
26.8889 | -106.2631
26.8925 | -106.2576
26.8962 | -106.2536
26.8999 | -106.249
26.9035 | -106.2454
26.9072 | -106.2407
26.9109 | -106.2351
26.9145 | -106.2311
26.9182 | -106.2275
26.9218 | -106.2249
26.9255 | -106.2222
26.9292 | -106.2183
26.9328 | -106.2156
26.9365 | -106.2124
26.9402 | -106.208
26.9438 | -106.2037
26.9475 | -106.2035
26.9511 | -106.2024
26.9548 | -106.2007
26.9585 | -106.1988
26.9621 | -106.1965
26.9658 | -106.1947
26.9695 | -106.1931
26.9731 | -106.1911
26.9768 | -106.1893
26.9805 | -106.1882
26.9841 | -106.1885
26.9878 | -106.1895
26.9914 | -106.1895
26.9951 | -106.1893
26.9988 | -106.1894
27.0024 | -106.1875
27.0061 | -106.1861
27.0098 | -106.1868
27.0134 | -106.1871
27.0171 | -106.188
27.0207 | -106.1864
27.0244 | -106.1849
27.0281 | -106.1845
27.0317 | -106.1833
27.0354 | -106.1843
27.0391 | -106.1853
27.0427 | -106.1864
27.0464 | -106.1873
27.05 | -106.1875
27.0537 | -106.1884
27.0574 | -106.1892
27.061 | -106.1888
27.0647 | -106.1881
27.0684 | -106.1881
27.072 | -106.1893
27.0757 | -106.1906
27.0793 | -106.1892
27.083 | -106.1884
27.0867 | -106.1879
27.0903 | -106.1887
27.094 | -106.1895
27.0977 | -106.1877
27.1013 | -106.1873
27.105 | -106.1866
27.1086 | -106.1855
27.1123 | -106.186
27.116 | -106.1852
27.1196 | -106.1837
27.1233 | -106.1817
27.127 | -106.1788
27.1306 | -106.1776
27.1343 | -106.1751
27.138 | -106.174
27.1416 | -106.1725
27.1453 | -106.1716
27.1489 | -106.1718
27.1526 | -106.1709
27.1563 | -106.1712
27.1599 | -106.171
27.1636 | -106.169
27.1673 | -106.1679
27.1709 | -106.1662
27.1746 | -106.1648
27.1782 | -106.1644
27.1819 | -106.1635
27.1856 | -106.1632
27.1892 | -106.1632
27.1929 | -106.1611
27.1966 | -106.1605
27.2002 | -106.1612
27.2039 | -106.1606
27.2075 | -106.1606
27.2112 | -106.1612
27.2149 | -106.161
27.2185 | -106.1603
27.2222 | -106.1577
27.2259 | -106.1567
27.2295 | -106.1558
27.2332 | -106.1552
27.2368 | -106.1551
27.2405 | -106.1552
27.2442 | -106.1564
27.2478 | -106.1576
27.2515 | -106.158
27.2552 | -106.1603
27.2588 | -106.1619
27.2625 | -106.1601
27.2662 | -106.1582
27.2698 | -106.1602
27.2735 | -106.1622
27.2771 | -106.1646
27.2808 | -106.1675
27.2845 | -106.1687
27.2881 | -106.169
27.2918 | -106.1714
27.2955 | -106.1731
27.2991 | -106.1736
27.3028 | -106.1738
27.3064 | -106.1753
27.3101 | -106.1763
27.3138 | -106.1757
27.3174 | -106.176
27.3211 | -106.1756
27.3248 | -106.1752
27.3284 | -106.1738
27.3321 | -106.1729
27.3357 | -106.1718
27.3394 | -106.1704
27.3431 | -106.1686
27.3467 | -106.1677
27.3504 | -106.1667
27.3541 | -106.1655
27.3577 | -106.1642
27.3614 | -106.1621
27.365 | -106.1594
27.3687 | -106.1549
27.3724 | -106.151
27.376 | -106.1492
27.3797 | -106.1483
27.3834 | -106.1479
27.387 | -106.148
27.3907 | -106.1466
27.3943 | -106.1437
27.398 | -106.1415
27.4017 | -106.1393
27.4053 | -106.1378
27.409 | -106.1362
27.4127 | -106.1355
27.4163 | -106.1349
27.42 | -106.1345
27.4237 | -106.1338
27.4273 | -106.1313
27.431 | -106.1302
27.4346 | -106.1321
27.4383 | -106.1327
27.442 | -106.1342
27.4456 | -106.1372
27.4493 | -106.1406
27.453 | -106.142
27.4566 | -106.1445
27.4603 | -106.1478
27.4639 | -106.1537
27.4676 | -106.1605
27.4713 | -106.1656
27.4749 | -106.1745
27.4786 | -106.1796
27.4823 | -106.1849
27.4859 | -106.1912
27.4896 | -106.197
27.4932 | -106.2047
27.4969 | -106.211
27.5006 | -106.2177
27.5042 | -106.2251
27.5079 | -106.2325
27.5116 | -106.2422
27.5152 | -106.251
27.5189 | -106.2582
27.5225 | -106.2659
27.5262 | -106.2732
27.5299 | -106.2814
27.5335 | -106.2883
27.5372 | -106.294
27.5409 | -106.3011
27.5445 | -106.3067
27.5482 | -106.3119
27.5519 | -106.3184
27.5555 | -106.324
27.5592 | -106.3302
27.5628 | -106.337
27.5665 | -106.3428
27.5702 | -106.3485
27.5738 | -106.3546
27.5775 | -106.3589
27.5812 | -106.3627
27.5848 | -106.3654
27.5885 | -106.369
27.5921 | -106.3726
27.5958 | -106.3748
27.5995 | -106.3777
27.6031 | -106.3825
27.6068 | -106.3859
27.6105 | -106.3876
27.6141 | -106.3875
27.6178 | -106.3881
27.6214 | -106.3873
27.6251 | -106.387
27.6288 | -106.3864
27.6324 | -106.3835
27.6361 | -106.381
27.6398 | -106.3794
27.6434 | -106.3768
27.6471 | -106.3742
27.6507 | -106.3722
27.6544 | -106.3691
27.6581 | -106.3633
27.6617 | -106.3574
27.6654 | -106.3511
27.6691 | -106.3432
27.6727 | -106.3366
27.6764 | -106.3297
27.6801 | -106.3208
27.6837 | -106.3141
27.6874 | -106.3068
27.691 | -106.2985
27.6947 | -106.2897
27.6984 | -106.2799
27.702 | -106.2716
27.7057 | -106.2618
27.7094 | -106.2531
27.713 | -106.2457
27.7167 | -106.2375
27.7203 | -106.2281
27.724 | -106.2191
27.7277 | -106.2103
27.7313 | -106.202
27.735 | -106.1932
27.7387 | -106.1849
27.7423 | -106.1754
27.746 | -106.1659
27.7496 | -106.1571
27.7533 | -106.1498
27.757 | -106.1424
27.7606 | -106.1345
27.7643 | -106.1257
27.768 | -106.1177
27.7716 | -106.1108
27.7753 | -106.1037
27.7789 | -106.0988
27.7826 | -106.094
27.7863 | -106.0888
27.7899 | -106.0837
27.7936 | -106.0791
27.7973 | -106.0747
27.8009 | -106.0701
27.8046 | -106.0663
27.8083 | -106.0619
27.8119 | -106.0585
27.8156 | -106.0566
27.8192 | -106.0532
27.8229 | -106.0483
27.8266 | -106.0443
27.8302 | -106.0407
27.8339 | -106.0385
27.8376 | -106.0373
27.8412 | -106.0384
27.8449 | -106.037
27.8485 | -106.0352
27.8522 | -106.0328
27.8559 | -106.032
27.8595 | -106.0326
27.8632 | -106.0326
27.8669 | -106.0331
27.8705 | -106.0334
27.8742 | -106.0328
27.8778 | -106.0332
27.8815 | -106.0345
27.8852 | -106.0343
27.8888 | -106.0364
27.8925 | -106.0354
27.8962 | -106.0347
27.8998 | -106.0366
27.9035 | -106.0384
27.9071 | -106.041
27.9108 | -106.0432
27.9145 | -106.044
27.9181 | -106.0443
27.9218 | -106.0464
27.9255 | -106.0481
27.9291 | -106.0498
27.9328 | -106.0519
27.9365 | -106.0546
27.9401 | -106.0556
27.9438 | -106.0555
27.9474 | -106.0568
27.9511 | -106.0583
27.9548 | -106.0592
27.9584 | -106.0585
27.9621 | -106.0584
27.9658 | -106.0581
27.9694 | -106.0582
27.9731 | -106.0592
27.9767 | -106.0594
27.9804 | -106.0576
27.9841 | -106.0559
27.9877 | -106.0545
27.9914 | -106.0503
27.9951 | -106.0467
27.9987 | -106.0443
28.0024 | -106.0429