## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 210 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 207
Hood | 02-92B/P | 147
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 987 | hPa |  | PRC PRO 01
Humidity | 39.3 | %/Hr
Internal temperature of calibration bench | 23.84 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 992 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.89 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.01 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3140 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 04/04/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 12/04/2018
Engineer | Anthony HUE | Approval |  | 23/04/2018
CEO | Robert GLEMOT | Approval |  | 23/04/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5203 | 57.5444
0.012209 | -36.0899 | 38.3581
0.018313 | -35.0283 | 28.2189
0.024417 | -34.565 | 21.975
0.030521 | -34.4532 | 18.216
0.036626 | -34.3082 | 15.1639
0.04273 | -34.2122 | 12.7978
0.048834 | -34.1665 | 11.2778
0.054938 | -34.1013 | 10.1422
0.061043 | -34.04 | 9.0603
0.067147 | -34.0294 | 8.2415
0.073251 | -34.0301 | 7.484
0.079355 | -34.0484 | 6.8227
0.08546 | -34.0625 | 6.2277
0.091564 | -34.0577 | 5.7878
0.097668 | -34.0393 | 5.5403
0.10377 | -34.0166 | 5.2908
0.10988 | -33.9988 | 5.0125
0.11598 | -33.9893 | 4.7196
0.12209 | -33.991 | 4.4369
0.12819 | -34.0009 | 4.1544
0.13429 | -34.0119 | 3.8665
0.1404 | -34.0154 | 3.5877
0.1465 | -34.0175 | 3.3297
0.15261 | -34.0149 | 3.1093
0.15871 | -34.0084 | 2.947
0.16482 | -34.0002 | 2.8443
0.17092 | -33.9977 | 2.8055
0.17702 | -33.9954 | 2.7336
0.18313 | -33.9942 | 2.6558
0.18923 | -33.9936 | 2.5471
0.19534 | -33.9921 | 2.4207
0.20144 | -33.9905 | 2.2786
0.20754 | -33.9874 | 2.1576
0.21365 | -33.9863 | 2.043
0.21975 | -33.9856 | 1.9451
0.22586 | -33.9866 | 1.858
0.23196 | -33.989 | 1.7861
0.23807 | -33.9931 | 1.7239
0.24417 | -33.9972 | 1.6656
0.25027 | -34.0011 | 1.6035
0.25638 | -34.0035 | 1.5417
0.26248 | -34.0053 | 1.4749
0.26859 | -34.0061 | 1.4095
0.27469 | -34.0064 | 1.3381
0.2808 | -34.0073 | 1.2641
0.2869 | -34.008 | 1.193
0.293 | -34.0088 | 1.1257
0.29911 | -34.0096 | 1.0655
0.30521 | -34.0105 | 1.0062
0.31132 | -34.0111 | 0.96742
0.31742 | -34.0123 | 0.92435
0.32353 | -34.0126 | 0.88923
0.32963 | -34.0134 | 0.86008
0.33573 | -34.0144 | 0.83971
0.34184 | -34.0157 | 0.81828
0.34794 | -34.0179 | 0.7945
0.35405 | -34.02 | 0.77045
0.36015 | -34.0225 | 0.74512
0.36626 | -34.0246 | 0.71756
0.37236 | -34.0268 | 0.68078
0.37846 | -34.029 | 0.6363
0.38457 | -34.0298 | 0.57865
0.39067 | -34.0309 | 0.53168
0.39678 | -34.0315 | 0.47934
0.40288 | -34.0319 | 0.42486
0.40899 | -34.0318 | 0.37249
0.41509 | -34.031 | 0.32116
0.42119 | -34.0309 | 0.27758
0.4273 | -34.0307 | 0.23097
0.4334 | -34.0306 | 0.18456
0.43951 | -34.0308 | 0.141
0.44561 | -34.0315 | 0.10108
0.45172 | -34.033 | 0.069356
0.45782 | -34.0337 | 0.032161
0.46392 | -34.0349 | -0.0041568
0.47003 | -34.0359 | -0.036577
0.47613 | -34.0367 | -0.070803
0.48224 | -34.0372 | -0.10575
0.48834 | -34.0362 | -0.14082
0.49445 | -34.0359 | -0.1763
0.50055 | -34.0353 | -0.20861
0.50665 | -34.0346 | -0.24311
0.51276 | -34.0335 | -0.27749
0.51886 | -34.0324 | -0.31179
0.52497 | -34.0312 | -0.33627
0.53107 | -34.0304 | -0.36608
0.53717 | -34.0297 | -0.39458
0.54328 | -34.0288 | -0.41771
0.54938 | -34.0286 | -0.44047
0.55549 | -34.0286 | -0.46163
0.56159 | -34.0297 | -0.47346
0.5677 | -34.03 | -0.4883
0.5738 | -34.031 | -0.50486
0.5799 | -34.0318 | -0.52006
0.58601 | -34.0329 | -0.53041
0.59211 | -34.0346 | -0.53688
0.59822 | -34.0359 | -0.54657
0.60432 | -34.0374 | -0.55888
0.61043 | -34.0387 | -0.567
0.61653 | -34.0399 | -0.5775
0.62263 | -34.0415 | -0.58708
0.62874 | -34.0421 | -0.59796
0.63484 | -34.0429 | -0.60745
0.64095 | -34.0434 | -0.61791
0.64705 | -34.0438 | -0.62789
0.65316 | -34.0439 | -0.64059
0.65926 | -34.044 | -0.65665
0.66536 | -34.0429 | -0.67436
0.67147 | -34.0425 | -0.68782
0.67757 | -34.0421 | -0.70241
0.68368 | -34.0413 | -0.72164
0.68978 | -34.0405 | -0.73918
0.69589 | -34.0395 | -0.75853
0.70199 | -34.0385 | -0.78165
0.70809 | -34.0376 | -0.80115
0.7142 | -34.0368 | -0.82197
0.7203 | -34.0357 | -0.84153
0.72641 | -34.035 | -0.86441
0.73251 | -34.0343 | -0.88477
0.73862 | -34.0345 | -0.90695
0.74472 | -34.0342 | -0.9273
0.75082 | -34.0339 | -0.9498
0.75693 | -34.0342 | -0.97201
0.76303 | -34.0347 | -0.99439
0.76914 | -34.0365 | -1.0152
0.77524 | -34.0374 | -1.0352
0.78135 | -34.0387 | -1.0556
0.78745 | -34.0401 | -1.0759
0.79355 | -34.0413 | -1.096
0.79966 | -34.043 | -1.1158
0.80576 | -34.0453 | -1.1358
0.81187 | -34.0469 | -1.1577
0.81797 | -34.0488 | -1.1793
0.82408 | -34.0505 | -1.2001
0.83018 | -34.0523 | -1.221
0.83628 | -34.0542 | -1.2397
0.84239 | -34.0553 | -1.2605
0.84849 | -34.0568 | -1.2825
0.8546 | -34.0582 | -1.3026
0.8607 | -34.0595 | -1.3237
0.86681 | -34.0606 | -1.3429
0.87291 | -34.0616 | -1.362
0.87901 | -34.062 | -1.3803
0.88512 | -34.0628 | -1.3986
0.89122 | -34.0635 | -1.4172
0.89733 | -34.0639 | -1.433
0.90343 | -34.0643 | -1.4508
0.90953 | -34.0645 | -1.4654
0.91564 | -34.0645 | -1.4815
0.92174 | -34.0647 | -1.4967
0.92785 | -34.0648 | -1.5116
0.93395 | -34.0649 | -1.5256
0.94006 | -34.0651 | -1.5387
0.94616 | -34.0648 | -1.5508
0.95226 | -34.0649 | -1.5637
0.95837 | -34.065 | -1.5778
0.96447 | -34.0652 | -1.5908
0.97058 | -34.065 | -1.6044
0.97668 | -34.0648 | -1.618
0.98279 | -34.0645 | -1.6309
0.98889 | -34.0643 | -1.6418
0.99499 | -34.0643 | -1.6553
1.0011 | -34.0641 | -1.6662
1.0072 | -34.0638 | -1.678
1.0133 | -34.0634 | -1.6889
1.0194 | -34.063 | -1.7021
1.0255 | -34.0626 | -1.716
1.0316 | -34.0621 | -1.7281
1.0377 | -34.0617 | -1.742
1.0438 | -34.0611 | -1.7549
1.0499 | -34.0603 | -1.7704
1.056 | -34.0596 | -1.7849
1.0621 | -34.059 | -1.7994
1.0682 | -34.0585 | -1.812
1.0743 | -34.0578 | -1.8277
1.0805 | -34.0573 | -1.8423
1.0866 | -34.0563 | -1.8576
1.0927 | -34.0555 | -1.8723
1.0988 | -34.0548 | -1.8874
1.1049 | -34.054 | -1.9017
1.111 | -34.0532 | -1.9173
1.1171 | -34.0525 | -1.9321
1.1232 | -34.052 | -1.9474
1.1293 | -34.0515 | -1.9616
1.1354 | -34.0512 | -1.9768
1.1415 | -34.0509 | -1.9925
1.1476 | -34.0502 | -2.0082
1.1537 | -34.0497 | -2.0231
1.1598 | -34.0498 | -2.038
1.1659 | -34.0495 | -2.0514
1.172 | -34.0493 | -2.0655
1.1781 | -34.049 | -2.0796
1.1842 | -34.0489 | -2.0939
1.1903 | -34.0492 | -2.1051
1.1964 | -34.0491 | -2.1186
1.2025 | -34.0492 | -2.1327
1.2086 | -34.0495 | -2.1452
1.2147 | -34.0496 | -2.1579
1.2209 | -34.0499 | -2.1692
1.227 | -34.0506 | -2.18
1.2331 | -34.0511 | -2.1921
1.2392 | -34.0514 | -2.2041
1.2453 | -34.0519 | -2.2153
1.2514 | -34.0523 | -2.2253
1.2575 | -34.0529 | -2.2364
1.2636 | -34.0537 | -2.2454
1.2697 | -34.0543 | -2.2548
1.2758 | -34.055 | -2.2661
1.2819 | -34.0557 | -2.2761
1.288 | -34.0564 | -2.2871
1.2941 | -34.0572 | -2.2956
1.3002 | -34.0579 | -2.3038
1.3063 | -34.0585 | -2.313
1.3124 | -34.0591 | -2.3221
1.3185 | -34.0599 | -2.3313
1.3246 | -34.0606 | -2.3405
1.3307 | -34.0614 | -2.3492
1.3368 | -34.0623 | -2.3581
1.3429 | -34.0629 | -2.3671
1.349 | -34.0634 | -2.3766
1.3551 | -34.0641 | -2.3861
1.3613 | -34.0647 | -2.3945
1.3674 | -34.0653 | -2.4046
1.3735 | -34.0657 | -2.4145
1.3796 | -34.0662 | -2.4237
1.3857 | -34.0665 | -2.4331
1.3918 | -34.0671 | -2.4434
1.3979 | -34.0676 | -2.4526
1.404 | -34.0681 | -2.4641
1.4101 | -34.0685 | -2.4746
1.4162 | -34.0689 | -2.4853
1.4223 | -34.0694 | -2.4968
1.4284 | -34.0699 | -2.5069
1.4345 | -34.0702 | -2.5186
1.4406 | -34.0704 | -2.5322
1.4467 | -34.0709 | -2.5439
1.4528 | -34.0713 | -2.5554
1.4589 | -34.0718 | -2.5674
1.465 | -34.0721 | -2.5784
1.4711 | -34.0726 | -2.5898
1.4772 | -34.0731 | -2.6033
1.4833 | -34.0735 | -2.6164
1.4894 | -34.074 | -2.6288
1.4955 | -34.0746 | -2.6408
1.5016 | -34.0751 | -2.6535
1.5078 | -34.0755 | -2.6675
1.5139 | -34.0759 | -2.6807
1.52 | -34.0763 | -2.6934
1.5261 | -34.0766 | -2.7062
1.5322 | -34.077 | -2.7204
1.5383 | -34.0773 | -2.7331
1.5444 | -34.0779 | -2.7459
1.5505 | -34.0783 | -2.7595
1.5566 | -34.0789 | -2.7726
1.5627 | -34.0793 | -2.7865
1.5688 | -34.0799 | -2.8003
1.5749 | -34.0803 | -2.813
1.581 | -34.0806 | -2.8266
1.5871 | -34.081 | -2.8405
1.5932 | -34.0813 | -2.8535
1.5993 | -34.0817 | -2.8679
1.6054 | -34.0821 | -2.8813
1.6115 | -34.0825 | -2.8956
1.6176 | -34.0828 | -2.9083
1.6237 | -34.083 | -2.9221
1.6298 | -34.0833 | -2.9348
1.6359 | -34.0836 | -2.9477
1.642 | -34.084 | -2.9598
1.6482 | -34.0842 | -2.9729
1.6543 | -34.0846 | -2.9868
1.6604 | -34.0849 | -2.999
1.6665 | -34.085 | -3.0116
1.6726 | -34.0852 | -3.0242
1.6787 | -34.0855 | -3.0369
1.6848 | -34.0854 | -3.0496
1.6909 | -34.0857 | -3.0625
1.697 | -34.0859 | -3.0748
1.7031 | -34.0859 | -3.0876
1.7092 | -34.0859 | -3.0998
1.7153 | -34.086 | -3.112
1.7214 | -34.0858 | -3.1243
1.7275 | -34.0858 | -3.1371
1.7336 | -34.0857 | -3.1502
1.7397 | -34.0858 | -3.1627
1.7458 | -34.0858 | -3.175
1.7519 | -34.0858 | -3.1874
1.758 | -34.0858 | -3.1995
1.7641 | -34.086 | -3.2112
1.7702 | -34.0859 | -3.2243
1.7763 | -34.0858 | -3.236
1.7824 | -34.0858 | -3.2485
1.7885 | -34.0858 | -3.2612
1.7947 | -34.0859 | -3.2734
1.8008 | -34.0858 | -3.2848
1.8069 | -34.0858 | -3.2965
1.813 | -34.0859 | -3.3088
1.8191 | -34.0858 | -3.3218
1.8252 | -34.0859 | -3.3347
1.8313 | -34.086 | -3.3471
1.8374 | -34.0861 | -3.3592
1.8435 | -34.0862 | -3.3713
1.8496 | -34.0862 | -3.3845
1.8557 | -34.0864 | -3.3975
1.8618 | -34.0866 | -3.4099
1.8679 | -34.0867 | -3.4221
1.874 | -34.0868 | -3.4343
1.8801 | -34.0869 | -3.4469
1.8862 | -34.0872 | -3.4591
1.8923 | -34.0873 | -3.4715
1.8984 | -34.0874 | -3.4838
1.9045 | -34.0874 | -3.4955
1.9106 | -34.0875 | -3.508
1.9167 | -34.0875 | -3.5204
1.9228 | -34.0876 | -3.5324
1.9289 | -34.0877 | -3.5445
1.9351 | -34.0879 | -3.5556
1.9412 | -34.0881 | -3.5671
1.9473 | -34.0884 | -3.5793
1.9534 | -34.0886 | -3.591
1.9595 | -34.0887 | -3.603
1.9656 | -34.089 | -3.6156
1.9717 | -34.0892 | -3.6275
1.9778 | -34.0895 | -3.6391
1.9839 | -34.0898 | -3.6508
1.99 | -34.0899 | -3.6621
1.9961 | -34.09 | -3.6745
2.0022 | -34.0903 | -3.6861
2.0083 | -34.0904 | -3.6975
2.0144 | -34.0906 | -3.7091
2.0205 | -34.0909 | -3.7211
2.0266 | -34.0912 | -3.7327
2.0327 | -34.0914 | -3.7445
2.0388 | -34.0919 | -3.7563
2.0449 | -34.0921 | -3.768
2.051 | -34.0923 | -3.7795
2.0571 | -34.0926 | -3.7911
2.0632 | -34.0929 | -3.8028
2.0693 | -34.093 | -3.8151
2.0754 | -34.0933 | -3.8267
2.0816 | -34.0936 | -3.839
2.0877 | -34.0938 | -3.8508
2.0938 | -34.0942 | -3.8631
2.0999 | -34.0944 | -3.875
2.106 | -34.0946 | -3.8868
2.1121 | -34.0948 | -3.899
2.1182 | -34.0951 | -3.91
2.1243 | -34.0953 | -3.9209
2.1304 | -34.0955 | -3.9323
2.1365 | -34.0958 | -3.9442
2.1426 | -34.0961 | -3.9564
2.1487 | -34.0964 | -3.9679
2.1548 | -34.0968 | -3.9799
2.1609 | -34.0973 | -3.9915
2.167 | -34.0977 | -4.0029
2.1731 | -34.098 | -4.0151
2.1792 | -34.0982 | -4.0264
2.1853 | -34.0986 | -4.0385
2.1914 | -34.0989 | -4.0501
2.1975 | -34.0993 | -4.0619
2.2036 | -34.0995 | -4.0742
2.2097 | -34.0998 | -4.0859
2.2158 | -34.1002 | -4.0967
2.222 | -34.1004 | -4.1074
2.2281 | -34.1007 | -4.119
2.2342 | -34.101 | -4.1306
2.2403 | -34.1012 | -4.1414
2.2464 | -34.1015 | -4.1524
2.2525 | -34.1018 | -4.164
2.2586 | -34.1021 | -4.1748
2.2647 | -34.1024 | -4.1854
2.2708 | -34.1028 | -4.1968
2.2769 | -34.1031 | -4.2077
2.283 | -34.1033 | -4.2191
2.2891 | -34.1035 | -4.2302
2.2952 | -34.1037 | -4.2409
2.3013 | -34.104 | -4.252
2.3074 | -34.1042 | -4.2627
2.3135 | -34.1044 | -4.2743
2.3196 | -34.1047 | -4.2847
2.3257 | -34.1049 | -4.2956
2.3318 | -34.105 | -4.3069
2.3379 | -34.1052 | -4.3179
2.344 | -34.1054 | -4.3295
2.3501 | -34.1056 | -4.3404
2.3562 | -34.1057 | -4.3509
2.3623 | -34.1059 | -4.3615
2.3685 | -34.106 | -4.3722
2.3746 | -34.1061 | -4.3836
2.3807 | -34.1064 | -4.3943
2.3868 | -34.1065 | -4.4052
2.3929 | -34.1067 | -4.4161
2.399 | -34.1068 | -4.4268
2.4051 | -34.1071 | -4.4375
2.4112 | -34.1073 | -4.4483
2.4173 | -34.1074 | -4.4591
2.4234 | -34.1075 | -4.4701
2.4295 | -34.1077 | -4.4811
2.4356 | -34.1078 | -4.4922
2.4417 | -34.108 | -4.5035
2.4478 | -34.1081 | -4.5145
2.4539 | -34.1082 | -4.5249
2.46 | -34.1083 | -4.5365
2.4661 | -34.1085 | -4.5478
2.4722 | -34.1086 | -4.5581
2.4783 | -34.1087 | -4.5687
2.4844 | -34.1088 | -4.5799
2.4905 | -34.1089 | -4.5911
2.4966 | -34.1091 | -4.602
2.5027 | -34.1091 | -4.6126
2.5089 | -34.1093 | -4.6236
2.515 | -34.1093 | -4.6346
2.5211 | -34.1094 | -4.6455
2.5272 | -34.1094 | -4.6565
2.5333 | -34.1096 | -4.6677
2.5394 | -34.1097 | -4.6786
2.5455 | -34.1098 | -4.6901
2.5516 | -34.11 | -4.7006
2.5577 | -34.1102 | -4.7121
2.5638 | -34.1103 | -4.724
2.5699 | -34.1103 | -4.7355
2.576 | -34.1104 | -4.7472
2.5821 | -34.1105 | -4.7586
2.5882 | -34.1106 | -4.7698
2.5943 | -34.1107 | -4.7808
2.6004 | -34.1109 | -4.7919
2.6065 | -34.111 | -4.8024
2.6126 | -34.1111 | -4.813
2.6187 | -34.1113 | -4.8235
2.6248 | -34.1114 | -4.8344
2.6309 | -34.1116 | -4.846
2.637 | -34.1118 | -4.8577
2.6431 | -34.1119 | -4.8683
2.6492 | -34.112 | -4.8797
2.6554 | -34.1122 | -4.8907
2.6615 | -34.1123 | -4.9018
2.6676 | -34.1126 | -4.9133
2.6737 | -34.1128 | -4.9246
2.6798 | -34.1131 | -4.9358
2.6859 | -34.1133 | -4.9472
2.692 | -34.1134 | -4.9582
2.6981 | -34.1136 | -4.9694
2.7042 | -34.1138 | -4.9807
2.7103 | -34.114 | -4.9914
2.7164 | -34.1143 | -5.0025
2.7225 | -34.1146 | -5.0139
2.7286 | -34.1148 | -5.0251
2.7347 | -34.115 | -5.0357
2.7408 | -34.1153 | -5.0459
2.7469 | -34.1155 | -5.0574
2.753 | -34.1158 | -5.0681
2.7591 | -34.1161 | -5.0795
2.7652 | -34.1163 | -5.0911
2.7713 | -34.1164 | -5.1019
2.7774 | -34.1167 | -5.113
2.7835 | -34.1169 | -5.1241
2.7896 | -34.1171 | -5.1347
2.7958 | -34.1174 | -5.1456
2.8019 | -34.1177 | -5.1561
2.808 | -34.1179 | -5.167
2.8141 | -34.1183 | -5.1783
2.8202 | -34.1186 | -5.1893
2.8263 | -34.1189 | -5.2003
2.8324 | -34.1192 | -5.2118
2.8385 | -34.1196 | -5.2228
2.8446 | -34.12 | -5.2337
2.8507 | -34.1204 | -5.2447
2.8568 | -34.1207 | -5.2553
2.8629 | -34.121 | -5.266
2.869 | -34.1213 | -5.2773
2.8751 | -34.1216 | -5.2882
2.8812 | -34.122 | -5.2997
2.8873 | -34.1223 | -5.3107
2.8934 | -34.1227 | -5.3219
2.8995 | -34.123 | -5.3331
2.9056 | -34.1233 | -5.3444
2.9117 | -34.1236 | -5.3556
2.9178 | -34.124 | -5.3665
2.9239 | -34.1244 | -5.3782
2.93 | -34.1248 | -5.3899
2.9361 | -34.1251 | -5.4009
2.9423 | -34.1255 | -5.4117
2.9484 | -34.1259 | -5.423
2.9545 | -34.1263 | -5.4342
2.9606 | -34.1267 | -5.4457
2.9667 | -34.127 | -5.4571
2.9728 | -34.1274 | -5.4686
2.9789 | -34.1278 | -5.4798
2.985 | -34.1283 | -5.4914
2.9911 | -34.1286 | -5.5029
2.9972 | -34.1291 | -5.5142
3.0033 | -34.1295 | -5.5259
3.0094 | -34.1299 | -5.5371
3.0155 | -34.1303 | -5.5483
3.0216 | -34.1307 | -5.5598
3.0277 | -34.131 | -5.5716
3.0338 | -34.1314 | -5.5833
3.0399 | -34.1319 | -5.5948
3.046 | -34.1322 | -5.6066
3.0521 | -34.1325 | -5.6179
3.0582 | -34.1328 | -5.6302
3.0643 | -34.1332 | -5.6419
3.0704 | -34.1335 | -5.654
3.0765 | -34.1338 | -5.6652
3.0827 | -34.1341 | -5.6767
3.0888 | -34.1345 | -5.6882
3.0949 | -34.1349 | -5.6998
3.101 | -34.1352 | -5.7115
3.1071 | -34.1355 | -5.7235
3.1132 | -34.1359 | -5.7355
3.1193 | -34.1363 | -5.7472
3.1254 | -34.1367 | -5.7588
3.1315 | -34.1371 | -5.771
3.1376 | -34.1373 | -5.7826
3.1437 | -34.1377 | -5.7944
3.1498 | -34.1381 | -5.8067
3.1559 | -34.1384 | -5.8184
3.162 | -34.1386 | -5.8307
3.1681 | -34.1389 | -5.8429
3.1742 | -34.1392 | -5.8548
3.1803 | -34.1396 | -5.8666
3.1864 | -34.14 | -5.8782
3.1925 | -34.1403 | -5.89
3.1986 | -34.1405 | -5.9022
3.2047 | -34.1408 | -5.9145
3.2108 | -34.1411 | -5.9259
3.2169 | -34.1414 | -5.9378
3.223 | -34.1417 | -5.9498
3.2292 | -34.1419 | -5.9619
3.2353 | -34.1421 | -5.973
3.2414 | -34.1424 | -5.9847
3.2475 | -34.1426 | -5.9963
3.2536 | -34.1429 | -6.008
3.2597 | -34.1432 | -6.0192
3.2658 | -34.1434 | -6.0304
3.2719 | -34.1436 | -6.0416
3.278 | -34.1439 | -6.0535
3.2841 | -34.1441 | -6.0648
3.2902 | -34.1444 | -6.0765
3.2963 | -34.1446 | -6.0882
3.3024 | -34.1447 | -6.0995
3.3085 | -34.145 | -6.1113
3.3146 | -34.1452 | -6.1227
3.3207 | -34.1455 | -6.1344
3.3268 | -34.1457 | -6.1461
3.3329 | -34.1459 | -6.1571
3.339 | -34.1461 | -6.1682
3.3451 | -34.1463 | -6.1796
3.3512 | -34.1464 | -6.1906
3.3573 | -34.1467 | -6.202
3.3634 | -34.1469 | -6.2134
3.3696 | -34.1471 | -6.225
3.3757 | -34.1472 | -6.236
3.3818 | -34.1475 | -6.2471
3.3879 | -34.1476 | -6.2584
3.394 | -34.1478 | -6.27
3.4001 | -34.148 | -6.281
3.4062 | -34.1481 | -6.2922
3.4123 | -34.1483 | -6.303
3.4184 | -34.1485 | -6.3145
3.4245 | -34.1487 | -6.3255
3.4306 | -34.1488 | -6.3364
3.4367 | -34.149 | -6.3471
3.4428 | -34.1491 | -6.3577
3.4489 | -34.1493 | -6.3689
3.455 | -34.1494 | -6.38
3.4611 | -34.1496 | -6.3914
3.4672 | -34.1498 | -6.4027
3.4733 | -34.15 | -6.4139
3.4794 | -34.1501 | -6.4249
3.4855 | -34.1503 | -6.4355
3.4916 | -34.1505 | -6.4461
3.4977 | -34.1506 | -6.4576
3.5038 | -34.1508 | -6.4687
3.5099 | -34.1509 | -6.4791
3.5161 | -34.1511 | -6.4897
3.5222 | -34.1512 | -6.5002
3.5283 | -34.1513 | -6.5112
3.5344 | -34.1514 | -6.5222
3.5405 | -34.1516 | -6.5329
3.5466 | -34.1518 | -6.544
3.5527 | -34.1519 | -6.555
3.5588 | -34.152 | -6.5658
3.5649 | -34.1521 | -6.5766
3.571 | -34.1522 | -6.5874
3.5771 | -34.1524 | -6.5987
3.5832 | -34.1525 | -6.6088
3.5893 | -34.1527 | -6.6196
3.5954 | -34.1529 | -6.6306
3.6015 | -34.1531 | -6.6417
3.6076 | -34.1533 | -6.6524
3.6137 | -34.1534 | -6.6627
3.6198 | -34.1537 | -6.6731
3.6259 | -34.1539 | -6.6842
3.632 | -34.154 | -6.695
3.6381 | -34.1542 | -6.7059
3.6442 | -34.1544 | -6.7167
3.6503 | -34.1546 | -6.727
3.6565 | -34.1548 | -6.7379
3.6626 | -34.1549 | -6.7487
3.6687 | -34.1552 | -6.7596
3.6748 | -34.1554 | -6.7702
3.6809 | -34.1557 | -6.781
3.687 | -34.1558 | -6.7922
3.6931 | -34.1561 | -6.8026
3.6992 | -34.1563 | -6.8134
3.7053 | -34.1565 | -6.824
3.7114 | -34.1567 | -6.8349
3.7175 | -34.1569 | -6.8457
3.7236 | -34.1572 | -6.856
3.7297 | -34.1574 | -6.8669
3.7358 | -34.1576 | -6.878
3.7419 | -34.1578 | -6.889
3.748 | -34.158 | -6.8998
3.7541 | -34.1582 | -6.9105
3.7602 | -34.1585 | -6.9212
3.7663 | -34.1587 | -6.9318
3.7724 | -34.159 | -6.9426
3.7785 | -34.1593 | -6.9532
3.7846 | -34.1595 | -6.9638
3.7907 | -34.1598 | -6.975
3.7969 | -34.1601 | -6.9859
3.803 | -34.1603 | -6.9962
3.8091 | -34.1606 | -7.007
3.8152 | -34.1608 | -7.0176
3.8213 | -34.1611 | -7.0286
3.8274 | -34.1613 | -7.0391
3.8335 | -34.1617 | -7.0501
3.8396 | -34.162 | -7.0609
3.8457 | -34.1623 | -7.0718
3.8518 | -34.1627 | -7.0825
3.8579 | -34.163 | -7.0932
3.864 | -34.1633 | -7.1042
3.8701 | -34.1636 | -7.1154
3.8762 | -34.1639 | -7.126
3.8823 | -34.1641 | -7.137
3.8884 | -34.1645 | -7.1478
3.8945 | -34.1648 | -7.1589
3.9006 | -34.1652 | -7.1701
3.9067 | -34.1654 | -7.1814
3.9128 | -34.1658 | -7.1923
3.9189 | -34.1661 | -7.2029
3.925 | -34.1663 | -7.2139
3.9311 | -34.1666 | -7.2243
3.9372 | -34.1669 | -7.2353
3.9434 | -34.1672 | -7.2462
3.9495 | -34.1676 | -7.2572
3.9556 | -34.168 | -7.2679
3.9617 | -34.1683 | -7.2784
3.9678 | -34.1686 | -7.2889
3.9739 | -34.1689 | -7.2998
3.98 | -34.1692 | -7.3107
3.9861 | -34.1695 | -7.3215
3.9922 | -34.1697 | -7.3324
3.9983 | -34.1701 | -7.3428
4.0044 | -34.1704 | -7.3532
4.0105 | -34.1707 | -7.3641
4.0166 | -34.171 | -7.375
4.0227 | -34.1713 | -7.3858
4.0288 | -34.1716 | -7.3971
4.0349 | -34.1718 | -7.4077
4.041 | -34.1721 | -7.4185
4.0471 | -34.1724 | -7.4294
4.0532 | -34.1727 | -7.4399
4.0593 | -34.173 | -7.4509
4.0654 | -34.1733 | -7.4615
4.0715 | -34.1736 | -7.4728
4.0776 | -34.174 | -7.4835
4.0838 | -34.1743 | -7.4948
4.0899 | -34.1746 | -7.5053
4.096 | -34.1749 | -7.5161
4.1021 | -34.1751 | -7.5267
4.1082 | -34.1754 | -7.5372
4.1143 | -34.1757 | -7.5477
4.1204 | -34.176 | -7.5583
4.1265 | -34.1763 | -7.5687
4.1326 | -34.1766 | -7.5791
4.1387 | -34.177 | -7.5898
4.1448 | -34.1772 | -7.6006
4.1509 | -34.1775 | -7.6113
4.157 | -34.1778 | -7.6225
4.1631 | -34.1781 | -7.6335
4.1692 | -34.1784 | -7.644
4.1753 | -34.1786 | -7.6547
4.1814 | -34.1789 | -7.6651
4.1875 | -34.1792 | -7.6761
4.1936 | -34.1795 | -7.6869
4.1997 | -34.1797 | -7.6973
4.2058 | -34.18 | -7.7083
4.2119 | -34.1802 | -7.7188
4.218 | -34.1805 | -7.7292
4.2241 | -34.1807 | -7.7397
4.2303 | -34.1811 | -7.7502
4.2364 | -34.1813 | -7.7609
4.2425 | -34.1815 | -7.7716
4.2486 | -34.1817 | -7.7824
4.2547 | -34.1819 | -7.7928
4.2608 | -34.1821 | -7.8036
4.2669 | -34.1823 | -7.8138
4.273 | -34.1826 | -7.8247
4.2791 | -34.1829 | -7.8353
4.2852 | -34.1831 | -7.8461
4.2913 | -34.1833 | -7.8564
4.2974 | -34.1835 | -7.8673
4.3035 | -34.1838 | -7.8782
4.3096 | -34.184 | -7.889
4.3157 | -34.1843 | -7.8995
4.3218 | -34.1845 | -7.91
4.3279 | -34.1847 | -7.9206
4.334 | -34.1849 | -7.9312
4.3401 | -34.1852 | -7.9413
4.3462 | -34.1855 | -7.952
4.3523 | -34.1857 | -7.9622
4.3584 | -34.1859 | -7.9723
4.3645 | -34.1861 | -7.9828
4.3707 | -34.1863 | -7.9931
4.3768 | -34.1866 | -8.0035
4.3829 | -34.1868 | -8.0145
4.389 | -34.187 | -8.025
4.3951 | -34.1873 | -8.0353
4.4012 | -34.1874 | -8.0462
4.4073 | -34.1876 | -8.0565
4.4134 | -34.1879 | -8.0669
4.4195 | -34.1882 | -8.0774
4.4256 | -34.1885 | -8.0877
4.4317 | -34.1887 | -8.0981
4.4378 | -34.1889 | -8.1082
4.4439 | -34.1891 | -8.1184
4.45 | -34.1892 | -8.1289
4.4561 | -34.1894 | -8.1393
4.4622 | -34.1896 | -8.1499
4.4683 | -34.1898 | -8.1603
4.4744 | -34.1901 | -8.171
4.4805 | -34.1903 | -8.1811
4.4866 | -34.1905 | -8.1915
4.4927 | -34.1907 | -8.2022
4.4988 | -34.1909 | -8.2126
4.5049 | -34.1912 | -8.2226
4.511 | -34.1914 | -8.2328
4.5172 | -34.1916 | -8.2431
4.5233 | -34.1918 | -8.2539
4.5294 | -34.1921 | -8.2642
4.5355 | -34.1923 | -8.2746
4.5416 | -34.1925 | -8.285
4.5477 | -34.1927 | -8.2954
4.5538 | -34.1929 | -8.3054
4.5599 | -34.1932 | -8.3157
4.566 | -34.1934 | -8.3261
4.5721 | -34.1936 | -8.3365
4.5782 | -34.1939 | -8.3471
4.5843 | -34.1941 | -8.3576
4.5904 | -34.1944 | -8.368
4.5965 | -34.1946 | -8.3783
4.6026 | -34.1948 | -8.3892
4.6087 | -34.195 | -8.3994
4.6148 | -34.1952 | -8.4102
4.6209 | -34.1955 | -8.4205
4.627 | -34.1957 | -8.4311
4.6331 | -34.1959 | -8.4414
4.6392 | -34.1961 | -8.4519
4.6453 | -34.1964 | -8.4622
4.6514 | -34.1966 | -8.4727
4.6576 | -34.1968 | -8.4835
4.6637 | -34.197 | -8.4935
4.6698 | -34.1972 | -8.5038
4.6759 | -34.1974 | -8.5146
4.682 | -34.1976 | -8.5248
4.6881 | -34.1978 | -8.5352
4.6942 | -34.198 | -8.5456
4.7003 | -34.1982 | -8.5561
4.7064 | -34.1985 | -8.5663
4.7125 | -34.1987 | -8.577
4.7186 | -34.1989 | -8.5871
4.7247 | -34.1992 | -8.5976
4.7308 | -34.1994 | -8.6079
4.7369 | -34.1996 | -8.6185
4.743 | -34.1999 | -8.6292
4.7491 | -34.2001 | -8.6399
4.7552 | -34.2003 | -8.6505
4.7613 | -34.2005 | -8.661
4.7674 | -34.2007 | -8.6713
4.7735 | -34.2009 | -8.6822
4.7796 | -34.2012 | -8.6929
4.7857 | -34.2014 | -8.7035
4.7918 | -34.2017 | -8.7137
4.7979 | -34.2019 | -8.7241
4.8041 | -34.2021 | -8.7346
4.8102 | -34.2024 | -8.7447
4.8163 | -34.2026 | -8.7551
4.8224 | -34.2028 | -8.7656
4.8285 | -34.2031 | -8.7757
4.8346 | -34.2034 | -8.786
4.8407 | -34.2035 | -8.7965
4.8468 | -34.2038 | -8.8072
4.8529 | -34.2041 | -8.818
4.859 | -34.2043 | -8.8288
4.8651 | -34.2045 | -8.8389
4.8712 | -34.2048 | -8.8492
4.8773 | -34.205 | -8.8597
4.8834 | -34.2052 | -8.8702
4.8895 | -34.2055 | -8.8805
4.8956 | -34.2057 | -8.8909
4.9017 | -34.2059 | -8.9012
4.9078 | -34.2061 | -8.912
4.9139 | -34.2064 | -8.922
4.92 | -34.2067 | -8.9322
4.9261 | -34.2069 | -8.9423
4.9322 | -34.2072 | -8.9527
4.9383 | -34.2075 | -8.9632
4.9445 | -34.2077 | -8.9736
4.9506 | -34.2079 | -8.9839
4.9567 | -34.2082 | -8.9941
4.9628 | -34.2084 | -9.0045
4.9689 | -34.2087 | -9.0153
4.975 | -34.2089 | -9.026
4.9811 | -34.2092 | -9.0362
4.9872 | -34.2095 | -9.0465
4.9933 | -34.2098 | -9.0573
4.9994 | -34.21 | -9.0681
5.0055 | -34.2103 | -9.0786
5.0116 | -34.2106 | -9.0895
5.0177 | -34.2109 | -9.1003
5.0238 | -34.2112 | -9.111
5.0299 | -34.2115 | -9.1219
5.036 | -34.2118 | -9.1324
5.0421 | -34.212 | -9.1428
5.0482 | -34.2123 | -9.1533
5.0543 | -34.2126 | -9.1635
5.0604 | -34.2129 | -9.174
5.0665 | -34.2132 | -9.1842
5.0726 | -34.2134 | -9.195
5.0787 | -34.2137 | -9.2057
5.0848 | -34.214 | -9.2163
5.091 | -34.2143 | -9.2266
5.0971 | -34.2146 | -9.2368
5.1032 | -34.2149 | -9.2476
5.1093 | -34.2152 | -9.2581
5.1154 | -34.2154 | -9.2684
5.1215 | -34.2156 | -9.2786
5.1276 | -34.2159 | -9.2889
5.1337 | -34.2163 | -9.2991
5.1398 | -34.2165 | -9.3094
5.1459 | -34.2168 | -9.3202
5.152 | -34.2171 | -9.3308
5.1581 | -34.2174 | -9.3415
5.1642 | -34.2177 | -9.3521
5.1703 | -34.218 | -9.3626
5.1764 | -34.2183 | -9.3729
5.1825 | -34.2186 | -9.3832
5.1886 | -34.2189 | -9.3934
5.1947 | -34.2191 | -9.4039
5.2008 | -34.2195 | -9.4145
5.2069 | -34.2197 | -9.4253
5.213 | -34.22 | -9.4356
5.2191 | -34.2204 | -9.4461
5.2252 | -34.2206 | -9.4567
5.2314 | -34.2209 | -9.4674
5.2375 | -34.2212 | -9.4776
5.2436 | -34.2215 | -9.4882
5.2497 | -34.2218 | -9.4985
5.2558 | -34.2221 | -9.5093
5.2619 | -34.2224 | -9.5201
5.268 | -34.2227 | -9.5308
5.2741 | -34.2231 | -9.5414
5.2802 | -34.2233 | -9.5524
5.2863 | -34.2237 | -9.5628
5.2924 | -34.224 | -9.5735
5.2985 | -34.2243 | -9.5841
5.3046 | -34.2246 | -9.5948
5.3107 | -34.2249 | -9.6055
5.3168 | -34.2253 | -9.6162
5.3229 | -34.2255 | -9.6268
5.329 | -34.2258 | -9.6373
5.3351 | -34.2262 | -9.6476
5.3412 | -34.2265 | -9.6581
5.3473 | -34.2268 | -9.6688
5.3534 | -34.2271 | -9.6795
5.3595 | -34.2274 | -9.6902
5.3656 | -34.2278 | -9.7007
5.3717 | -34.228 | -9.7117
5.3779 | -34.2283 | -9.7226
5.384 | -34.2287 | -9.7332
5.3901 | -34.229 | -9.744
5.3962 | -34.2293 | -9.7544
5.4023 | -34.2296 | -9.7653
5.4084 | -34.23 | -9.7758
5.4145 | -34.2303 | -9.7862
5.4206 | -34.2306 | -9.7968
5.4267 | -34.2309 | -9.8078
5.4328 | -34.2312 | -9.8184
5.4389 | -34.2316 | -9.8288
5.445 | -34.2319 | -9.8398
5.4511 | -34.2322 | -9.8502
5.4572 | -34.2325 | -9.8609
5.4633 | -34.2328 | -9.8715
5.4694 | -34.2331 | -9.8822
5.4755 | -34.2334 | -9.8925
5.4816 | -34.2338 | -9.9032
5.4877 | -34.2341 | -9.9137
5.4938 | -34.2344 | -9.9241
5.4999 | -34.2348 | -9.9346
5.506 | -34.2351 | -9.9452
5.5121 | -34.2354 | -9.9557
5.5183 | -34.2357 | -9.9662
5.5244 | -34.236 | -9.977
5.5305 | -34.2364 | -9.9876
5.5366 | -34.2367 | -9.9979
5.5427 | -34.237 | -10.0087
5.5488 | -34.2373 | -10.0194
5.5549 | -34.2376 | -10.0301
5.561 | -34.2379 | -10.0409
5.5671 | -34.2383 | -10.0514
5.5732 | -34.2386 | -10.0623
5.5793 | -34.2389 | -10.0733
5.5854 | -34.2393 | -10.0838
5.5915 | -34.2396 | -10.0948
5.5976 | -34.2399 | -10.1055
5.6037 | -34.2402 | -10.1162
5.6098 | -34.2405 | -10.1268
5.6159 | -34.2408 | -10.1373
5.622 | -34.2412 | -10.148
5.6281 | -34.2415 | -10.1588
5.6342 | -34.2418 | -10.1697
5.6403 | -34.2421 | -10.1804
5.6464 | -34.2425 | -10.1912
5.6525 | -34.2428 | -10.202
5.6586 | -34.2431 | -10.2123
5.6648 | -34.2435 | -10.223
5.6709 | -34.2439 | -10.2338
5.677 | -34.2442 | -10.2445
5.6831 | -34.2445 | -10.2553
5.6892 | -34.2448 | -10.2661
5.6953 | -34.2451 | -10.2768
5.7014 | -34.2455 | -10.2876
5.7075 | -34.2457 | -10.2982
5.7136 | -34.2461 | -10.3088
5.7197 | -34.2464 | -10.3195
5.7258 | -34.2467 | -10.3299
5.7319 | -34.247 | -10.3407
5.738 | -34.2474 | -10.3513
5.7441 | -34.2477 | -10.3619
5.7502 | -34.2481 | -10.3723
5.7563 | -34.2484 | -10.3829
5.7624 | -34.2488 | -10.3938
5.7685 | -34.2491 | -10.4045
5.7746 | -34.2494 | -10.4155
5.7807 | -34.2497 | -10.4262
5.7868 | -34.25 | -10.437
5.7929 | -34.2503 | -10.4476
5.799 | -34.2506 | -10.4581
5.8052 | -34.2509 | -10.4686
5.8113 | -34.2512 | -10.4792
5.8174 | -34.2516 | -10.4897
5.8235 | -34.2519 | -10.5003
5.8296 | -34.2522 | -10.5112
5.8357 | -34.2525 | -10.5218
5.8418 | -34.2528 | -10.5326
5.8479 | -34.2531 | -10.5433
5.854 | -34.2534 | -10.5541
5.8601 | -34.2538 | -10.5646
5.8662 | -34.2541 | -10.5754
5.8723 | -34.2544 | -10.5858
5.8784 | -34.2547 | -10.5964
5.8845 | -34.255 | -10.607
5.8906 | -34.2553 | -10.6177
5.8967 | -34.2556 | -10.6281
5.9028 | -34.256 | -10.6386
5.9089 | -34.2564 | -10.6491
5.915 | -34.2567 | -10.6595
5.9211 | -34.257 | -10.6701
5.9272 | -34.2573 | -10.6804
5.9333 | -34.2576 | -10.691
5.9394 | -34.258 | -10.7016
5.9455 | -34.2583 | -10.712
5.9517 | -34.2587 | -10.7224
5.9578 | -34.259 | -10.7331
5.9639 | -34.2593 | -10.7433
5.97 | -34.2597 | -10.754
5.9761 | -34.2601 | -10.7645
5.9822 | -34.2604 | -10.7752
5.9883 | -34.2607 | -10.7858
5.9944 | -34.261 | -10.7964
6.0005 | -34.2614 | -10.8072
6.0066 | -34.2617 | -10.8176
6.0127 | -34.2621 | -10.8279
6.0188 | -34.2624 | -10.8387
6.0249 | -34.2627 | -10.8494
6.031 | -34.263 | -10.8598
6.0371 | -34.2633 | -10.8704
6.0432 | -34.2636 | -10.8808
6.0493 | -34.264 | -10.8912
6.0554 | -34.2643 | -10.9016
6.0615 | -34.2647 | -10.9121
6.0676 | -34.265 | -10.9229
6.0737 | -34.2654 | -10.9335
6.0798 | -34.2657 | -10.9441
6.0859 | -34.2661 | -10.9545
6.0921 | -34.2664 | -10.9648
6.0982 | -34.2667 | -10.9753
6.1043 | -34.2671 | -10.9857
6.1104 | -34.2674 | -10.9965
6.1165 | -34.2678 | -11.0065
6.1226 | -34.2681 | -11.017
6.1287 | -34.2684 | -11.0275
6.1348 | -34.2688 | -11.0379
6.1409 | -34.2691 | -11.0484
6.147 | -34.2694 | -11.0588
6.1531 | -34.2697 | -11.0692
6.1592 | -34.2701 | -11.0796
6.1653 | -34.2704 | -11.09
6.1714 | -34.2707 | -11.1005
6.1775 | -34.2711 | -11.1111
6.1836 | -34.2714 | -11.1216
6.1897 | -34.2717 | -11.132
6.1958 | -34.2721 | -11.1428
6.2019 | -34.2724 | -11.1534
6.208 | -34.2728 | -11.1639
6.2141 | -34.2731 | -11.1743
6.2202 | -34.2735 | -11.1848
6.2263 | -34.2738 | -11.1952
6.2325 | -34.2741 | -11.2053
6.2386 | -34.2745 | -11.2155
6.2447 | -34.2748 | -11.2262
6.2508 | -34.2752 | -11.2367
6.2569 | -34.2755 | -11.247
6.263 | -34.2758 | -11.2575
6.2691 | -34.2762 | -11.2678
6.2752 | -34.2765 | -11.2784
6.2813 | -34.2768 | -11.2888
6.2874 | -34.2771 | -11.2991
6.2935 | -34.2775 | -11.3094
6.2996 | -34.2778 | -11.3198
6.3057 | -34.2781 | -11.3303
6.3118 | -34.2784 | -11.3409
6.3179 | -34.2787 | -11.3512
6.324 | -34.2791 | -11.3616
6.3301 | -34.2794 | -11.372
6.3362 | -34.2798 | -11.3823
6.3423 | -34.2801 | -11.3929
6.3484 | -34.2805 | -11.4033
6.3545 | -34.2808 | -11.4134
6.3606 | -34.2811 | -11.4236
6.3667 | -34.2814 | -11.4338
6.3728 | -34.2817 | -11.4442
6.379 | -34.2821 | -11.4546
6.3851 | -34.2824 | -11.4648
6.3912 | -34.2827 | -11.4751
6.3973 | -34.283 | -11.4855
6.4034 | -34.2833 | -11.4956
6.4095 | -34.2837 | -11.506
6.4156 | -34.284 | -11.5164
6.4217 | -34.2843 | -11.5267
6.4278 | -34.2846 | -11.5372
6.4339 | -34.285 | -11.5474
6.44 | -34.2853 | -11.5577
6.4461 | -34.2856 | -11.5683
6.4522 | -34.286 | -11.5784
6.4583 | -34.2863 | -11.5889
6.4644 | -34.2866 | -11.5993
6.4705 | -34.2869 | -11.6093
6.4766 | -34.2873 | -11.6197
6.4827 | -34.2876 | -11.6301
6.4888 | -34.2879 | -11.6402
6.4949 | -34.2882 | -11.6503
6.501 | -34.2885 | -11.6607
6.5071 | -34.2889 | -11.6708
6.5132 | -34.2892 | -11.6811
6.5194 | -34.2895 | -11.6915
6.5255 | -34.2899 | -11.7016
6.5316 | -34.2901 | -11.712
6.5377 | -34.2905 | -11.7223
6.5438 | -34.2908 | -11.7324
6.5499 | -34.2911 | -11.7426
6.556 | -34.2914 | -11.7528
6.5621 | -34.2917 | -11.7634
6.5682 | -34.2921 | -11.7737
6.5743 | -34.2925 | -11.7839
6.5804 | -34.2928 | -11.7943
6.5865 | -34.2931 | -11.8046
6.5926 | -34.2934 | -11.8148
6.5987 | -34.2938 | -11.8248
6.6048 | -34.2941 | -11.8351
6.6109 | -34.2944 | -11.8453
6.617 | -34.2947 | -11.8557
6.6231 | -34.295 | -11.8659
6.6292 | -34.2954 | -11.8762
6.6353 | -34.2957 | -11.8868
6.6414 | -34.2961 | -11.8969
6.6475 | -34.2964 | -11.9071
6.6536 | -34.2967 | -11.9173
6.6597 | -34.297 | -11.9272
6.6659 | -34.2973 | -11.9375
6.672 | -34.2977 | -11.9476
6.6781 | -34.298 | -11.9579
6.6842 | -34.2983 | -11.968
6.6903 | -34.2987 | -11.9782
6.6964 | -34.2991 | -11.9885
6.7025 | -34.2994 | -11.9985
6.7086 | -34.2997 | -12.0087
6.7147 | -34.3 | -12.019
6.7208 | -34.3003 | -12.0294
6.7269 | -34.3006 | -12.0395
6.733 | -34.3009 | -12.0495
6.7391 | -34.3013 | -12.0597
6.7452 | -34.3016 | -12.0697
6.7513 | -34.302 | -12.0796
6.7574 | -34.3023 | -12.0899
6.7635 | -34.3026 | -12.0999
6.7696 | -34.303 | -12.11
6.7757 | -34.3033 | -12.1202
6.7818 | -34.3037 | -12.1304
6.7879 | -34.304 | -12.1404
6.794 | -34.3043 | -12.1504
6.8001 | -34.3046 | -12.1605
6.8063 | -34.305 | -12.1706
6.8124 | -34.3053 | -12.1809
6.8185 | -34.3056 | -12.1913
6.8246 | -34.3059 | -12.2012
6.8307 | -34.3062 | -12.2112
6.8368 | -34.3066 | -12.2216
6.8429 | -34.3069 | -12.2318
6.849 | -34.3072 | -12.242
6.8551 | -34.3076 | -12.252
6.8612 | -34.3079 | -12.2621
6.8673 | -34.3082 | -12.2722
6.8734 | -34.3086 | -12.2825
6.8795 | -34.3089 | -12.2926
6.8856 | -34.3092 | -12.3028
6.8917 | -34.3096 | -12.3128
6.8978 | -34.3099 | -12.3227
6.9039 | -34.3102 | -12.3325
6.91 | -34.3106 | -12.3427
6.9161 | -34.3109 | -12.3526
6.9222 | -34.3112 | -12.3627
6.9283 | -34.3116 | -12.3725
6.9344 | -34.312 | -12.3827
6.9405 | -34.3123 | -12.3927
6.9466 | -34.3127 | -12.403
6.9528 | -34.313 | -12.4133
6.9589 | -34.3134 | -12.4236
6.965 | -34.3137 | -12.4337
6.9711 | -34.3141 | -12.4439
6.9772 | -34.3145 | -12.454
6.9833 | -34.3148 | -12.4641
6.9894 | -34.3151 | -12.4745
6.9955 | -34.3155 | -12.4843
7.0016 | -34.3158 | -12.4943
7.0077 | -34.3162 | -12.5044
7.0138 | -34.3165 | -12.5145
7.0199 | -34.3168 | -12.5244
7.026 | -34.3172 | -12.5345
7.0321 | -34.3176 | -12.5443
7.0382 | -34.3179 | -12.5541
7.0443 | -34.3183 | -12.5643
7.0504 | -34.3186 | -12.5745
7.0565 | -34.319 | -12.5846
7.0626 | -34.3193 | -12.5944
7.0687 | -34.3197 | -12.604
7.0748 | -34.32 | -12.6138
7.0809 | -34.3204 | -12.6238
7.087 | -34.3207 | -12.6339
7.0932 | -34.321 | -12.644
7.0993 | -34.3214 | -12.654
7.1054 | -34.3218 | -12.6639
7.1115 | -34.3221 | -12.6737
7.1176 | -34.3225 | -12.6837
7.1237 | -34.3228 | -12.6936
7.1298 | -34.3231 | -12.7035
7.1359 | -34.3235 | -12.7135
7.142 | -34.3238 | -12.7233
7.1481 | -34.3242 | -12.7332
7.1542 | -34.3245 | -12.7432
7.1603 | -34.3249 | -12.7532
7.1664 | -34.3252 | -12.7632
7.1725 | -34.3256 | -12.7732
7.1786 | -34.326 | -12.7832
7.1847 | -34.3263 | -12.7932
7.1908 | -34.3267 | -12.8033
7.1969 | -34.327 | -12.8135
7.203 | -34.3273 | -12.8233
7.2091 | -34.3277 | -12.8332
7.2152 | -34.3281 | -12.8432
7.2213 | -34.3284 | -12.8531
7.2274 | -34.3288 | -12.863
7.2335 | -34.3291 | -12.8729
7.2397 | -34.3295 | -12.8828
7.2458 | -34.3299 | -12.8926
7.2519 | -34.3303 | -12.9026
7.258 | -34.3306 | -12.9127
7.2641 | -34.331 | -12.9227
7.2702 | -34.3313 | -12.9328
7.2763 | -34.3317 | -12.9428
7.2824 | -34.3321 | -12.9529
7.2885 | -34.3324 | -12.9629
7.2946 | -34.3328 | -12.9729
7.3007 | -34.3331 | -12.983
7.3068 | -34.3334 | -12.9932
7.3129 | -34.3338 | -13.0033
7.319 | -34.3342 | -13.0133
7.3251 | -34.3345 | -13.0231
7.3312 | -34.3349 | -13.0331
7.3373 | -34.3353 | -13.0428
7.3434 | -34.3356 | -13.0527
7.3495 | -34.336 | -13.0625
7.3556 | -34.3364 | -13.0726
7.3617 | -34.3368 | -13.0827
7.3678 | -34.3371 | -13.0926
7.3739 | -34.3374 | -13.1028
7.3801 | -34.3378 | -13.113
7.3862 | -34.3382 | -13.1229
7.3923 | -34.3386 | -13.133
7.3984 | -34.3389 | -13.1427
7.4045 | -34.3393 | -13.1527
7.4106 | -34.3397 | -13.1626
7.4167 | -34.34 | -13.1726
7.4228 | -34.3404 | -13.1827
7.4289 | -34.3408 | -13.1926
7.435 | -34.3412 | -13.2024
7.4411 | -34.3415 | -13.2125
7.4472 | -34.3419 | -13.2225
7.4533 | -34.3423 | -13.2323
7.4594 | -34.3427 | -13.2425
7.4655 | -34.343 | -13.2523
7.4716 | -34.3434 | -13.2624
7.4777 | -34.3437 | -13.2723
7.4838 | -34.3441 | -13.2822
7.4899 | -34.3445 | -13.2921
7.496 | -34.3448 | -13.3019
7.5021 | -34.3452 | -13.3118
7.5082 | -34.3456 | -13.3218
7.5143 | -34.346 | -13.3318
7.5204 | -34.3463 | -13.3419
7.5266 | -34.3467 | -13.3517
7.5327 | -34.3471 | -13.3616
7.5388 | -34.3474 | -13.3719
7.5449 | -34.3478 | -13.382
7.551 | -34.3482 | -13.3918
7.5571 | -34.3486 | -13.4019
7.5632 | -34.349 | -13.412
7.5693 | -34.3494 | -13.4218
7.5754 | -34.3497 | -13.4317
7.5815 | -34.3501 | -13.4417
7.5876 | -34.3504 | -13.4519
7.5937 | -34.3509 | -13.462
7.5998 | -34.3513 | -13.4719
7.6059 | -34.3517 | -13.482
7.612 | -34.352 | -13.492
7.6181 | -34.3524 | -13.5019
7.6242 | -34.3528 | -13.5122
7.6303 | -34.3532 | -13.5224
7.6364 | -34.3535 | -13.5324
7.6425 | -34.3539 | -13.5424
7.6486 | -34.3542 | -13.5525
7.6547 | -34.3546 | -13.5623
7.6608 | -34.355 | -13.5723
7.667 | -34.3554 | -13.5824
7.6731 | -34.3557 | -13.5925
7.6792 | -34.3561 | -13.6027
7.6853 | -34.3565 | -13.6128
7.6914 | -34.3569 | -13.623
7.6975 | -34.3573 | -13.6331
7.7036 | -34.3577 | -13.6432
7.7097 | -34.358 | -13.6533
7.7158 | -34.3584 | -13.6634
7.7219 | -34.3588 | -13.6733
7.728 | -34.3592 | -13.6835
7.7341 | -34.3596 | -13.6936
7.7402 | -34.36 | -13.7037
7.7463 | -34.3604 | -13.7138
7.7524 | -34.3607 | -13.7241
7.7585 | -34.3611 | -13.7342
7.7646 | -34.3615 | -13.7442
7.7707 | -34.3619 | -13.7544
7.7768 | -34.3622 | -13.7646
7.7829 | -34.3626 | -13.775
7.789 | -34.363 | -13.7853
7.7951 | -34.3633 | -13.7953
7.8012 | -34.3637 | -13.8054
7.8073 | -34.3642 | -13.8157
7.8135 | -34.3645 | -13.8258
7.8196 | -34.3649 | -13.8361
7.8257 | -34.3653 | -13.8464
7.8318 | -34.3656 | -13.8565
7.8379 | -34.3661 | -13.8665
7.844 | -34.3664 | -13.8765
7.8501 | -34.3668 | -13.8868
7.8562 | -34.3672 | -13.8971
7.8623 | -34.3676 | -13.9071
7.8684 | -34.368 | -13.9173
7.8745 | -34.3683 | -13.9274
7.8806 | -34.3687 | -13.9377
7.8867 | -34.3691 | -13.9476
7.8928 | -34.3695 | -13.9575
7.8989 | -34.3699 | -13.9676
7.905 | -34.3703 | -13.9778
7.9111 | -34.3706 | -13.988
7.9172 | -34.371 | -13.9981
7.9233 | -34.3714 | -14.0083
7.9294 | -34.3718 | -14.0185
7.9355 | -34.3722 | -14.0284
7.9416 | -34.3725 | -14.0386
7.9477 | -34.3729 | -14.0485
7.9539 | -34.3733 | -14.0586
7.96 | -34.3737 | -14.0688
7.9661 | -34.3741 | -14.079
7.9722 | -34.3745 | -14.089
7.9783 | -34.3749 | -14.0994
7.9844 | -34.3753 | -14.1094
7.9905 | -34.3757 | -14.1198
7.9966 | -34.376 | -14.13
8.0027 | -34.3764 | -14.1401
8.0088 | -34.3768 | -14.1502
8.0149 | -34.3771 | -14.1604
8.021 | -34.3775 | -14.1706
8.0271 | -34.3779 | -14.1811
8.0332 | -34.3783 | -14.1911
8.0393 | -34.3787 | -14.2014
8.0454 | -34.3791 | -14.2118
8.0515 | -34.3795 | -14.222
8.0576 | -34.3799 | -14.2324
8.0637 | -34.3803 | -14.2424
8.0698 | -34.3806 | -14.2524
8.0759 | -34.381 | -14.2625
8.082 | -34.3814 | -14.2725
8.0881 | -34.3818 | -14.2828
8.0942 | -34.3822 | -14.293
8.1004 | -34.3826 | -14.3032
8.1065 | -34.383 | -14.3132
8.1126 | -34.3833 | -14.3234
8.1187 | -34.3837 | -14.3336
8.1248 | -34.3841 | -14.3436
8.1309 | -34.3845 | -14.3536
8.137 | -34.3849 | -14.3636
8.1431 | -34.3853 | -14.3739
8.1492 | -34.3856 | -14.384
8.1553 | -34.386 | -14.3942
8.1614 | -34.3864 | -14.4046
8.1675 | -34.3868 | -14.4147
8.1736 | -34.3871 | -14.425
8.1797 | -34.3875 | -14.4352
8.1858 | -34.3879 | -14.4455
8.1919 | -34.3883 | -14.4556
8.198 | -34.3887 | -14.4656
8.2041 | -34.3891 | -14.4758
8.2102 | -34.3895 | -14.4859
8.2163 | -34.3899 | -14.4959
8.2224 | -34.3902 | -14.506
8.2285 | -34.3906 | -14.5159
8.2346 | -34.391 | -14.526
8.2408 | -34.3914 | -14.536
8.2469 | -34.3918 | -14.5463
8.253 | -34.3922 | -14.5563
8.2591 | -34.3926 | -14.5665
8.2652 | -34.393 | -14.5767
8.2713 | -34.3934 | -14.5868
8.2774 | -34.3938 | -14.5971
8.2835 | -34.3942 | -14.6075
8.2896 | -34.3946 | -14.6175
8.2957 | -34.395 | -14.6277
8.3018 | -34.3953 | -14.6378
8.3079 | -34.3957 | -14.6478
8.314 | -34.3961 | -14.6579
8.3201 | -34.3966 | -14.668
8.3262 | -34.3969 | -14.6781
8.3323 | -34.3973 | -14.6882
8.3384 | -34.3977 | -14.6985
8.3445 | -34.3981 | -14.7086
8.3506 | -34.3985 | -14.7188
8.3567 | -34.3989 | -14.729
8.3628 | -34.3993 | -14.7391
8.3689 | -34.3997 | -14.7491
8.375 | -34.4 | -14.7593
8.3812 | -34.4004 | -14.7693
8.3873 | -34.4008 | -14.7795
8.3934 | -34.4012 | -14.7897
8.3995 | -34.4016 | -14.7997
8.4056 | -34.402 | -14.8098
8.4117 | -34.4023 | -14.82
8.4178 | -34.4027 | -14.8301
8.4239 | -34.4031 | -14.8401
8.43 | -34.4035 | -14.8502
8.4361 | -34.4038 | -14.8604
8.4422 | -34.4043 | -14.8704
8.4483 | -34.4046 | -14.8804
8.4544 | -34.4051 | -14.8906
8.4605 | -34.4054 | -14.9005
8.4666 | -34.4058 | -14.9109
8.4727 | -34.4062 | -14.921
8.4788 | -34.4066 | -14.931
8.4849 | -34.407 | -14.9411
8.491 | -34.4074 | -14.9513
8.4971 | -34.4078 | -14.9611
8.5032 | -34.4081 | -14.9711
8.5093 | -34.4085 | -14.981
8.5154 | -34.4089 | -14.9911
8.5215 | -34.4093 | -15.0011
8.5277 | -34.4097 | -15.0112
8.5338 | -34.4101 | -15.0212
8.5399 | -34.4104 | -15.0313
8.546 | -34.4109 | -15.0411
8.5521 | -34.4113 | -15.0512
8.5582 | -34.4116 | -15.0612
8.5643 | -34.412 | -15.071
8.5704 | -34.4124 | -15.0808
8.5765 | -34.4128 | -15.0909
8.5826 | -34.4132 | -15.1006
8.5887 | -34.4136 | -15.1106
8.5948 | -34.414 | -15.1207
8.6009 | -34.4144 | -15.1305
8.607 | -34.4148 | -15.1407
8.6131 | -34.4152 | -15.1508
8.6192 | -34.4156 | -15.1606
8.6253 | -34.416 | -15.1705
8.6314 | -34.4164 | -15.1805
8.6375 | -34.4167 | -15.1906
8.6436 | -34.4171 | -15.2006
8.6497 | -34.4175 | -15.2104
8.6558 | -34.4179 | -15.2204
8.6619 | -34.4182 | -15.2304
8.6681 | -34.4186 | -15.2406
8.6742 | -34.419 | -15.2505
8.6803 | -34.4194 | -15.2605
8.6864 | -34.4198 | -15.2702
8.6925 | -34.4202 | -15.28
8.6986 | -34.4206 | -15.29
8.7047 | -34.421 | -15.3002
8.7108 | -34.4213 | -15.3102
8.7169 | -34.4217 | -15.3202
8.723 | -34.4221 | -15.3299
8.7291 | -34.4225 | -15.3397
8.7352 | -34.4229 | -15.3499
8.7413 | -34.4233 | -15.3599
8.7474 | -34.4237 | -15.3697
8.7535 | -34.4241 | -15.3797
8.7596 | -34.4245 | -15.3896
8.7657 | -34.4249 | -15.3997
8.7718 | -34.4253 | -15.4097
8.7779 | -34.4257 | -15.4197
8.784 | -34.4261 | -15.4297
8.7901 | -34.4265 | -15.4396
8.7962 | -34.4269 | -15.4493
8.8023 | -34.4272 | -15.4592
8.8084 | -34.4276 | -15.4691
8.8146 | -34.428 | -15.4789
8.8207 | -34.4284 | -15.4888
8.8268 | -34.4288 | -15.4984
8.8329 | -34.4292 | -15.5083
8.839 | -34.4296 | -15.5182
8.8451 | -34.43 | -15.528
8.8512 | -34.4304 | -15.5379
8.8573 | -34.4308 | -15.5478
8.8634 | -34.4312 | -15.5574
8.8695 | -34.4316 | -15.5675
8.8756 | -34.4319 | -15.5774
8.8817 | -34.4323 | -15.5872
8.8878 | -34.4327 | -15.5973
8.8939 | -34.4331 | -15.6072
8.9 | -34.4335 | -15.617
8.9061 | -34.4339 | -15.6269
8.9122 | -34.4343 | -15.6367
8.9183 | -34.4347 | -15.6467
8.9244 | -34.4351 | -15.6567
8.9305 | -34.4355 | -15.6663
8.9366 | -34.4359 | -15.6761
8.9427 | -34.4363 | -15.6861
8.9488 | -34.4368 | -15.6959
8.955 | -34.4371 | -15.7057
8.9611 | -34.4375 | -15.7157
8.9672 | -34.438 | -15.7255
8.9733 | -34.4384 | -15.7352
8.9794 | -34.4388 | -15.7448
8.9855 | -34.4392 | -15.7548
8.9916 | -34.4396 | -15.7648
8.9977 | -34.44 | -15.7747
9.0038 | -34.4404 | -15.7844
9.0099 | -34.4408 | -15.7943
9.016 | -34.4412 | -15.8042
9.0221 | -34.4416 | -15.8138
9.0282 | -34.442 | -15.8235
9.0343 | -34.4424 | -15.8334
9.0404 | -34.4429 | -15.8432
9.0465 | -34.4433 | -15.8532
9.0526 | -34.4437 | -15.8629
9.0587 | -34.4442 | -15.8728
9.0648 | -34.4446 | -15.8825
9.0709 | -34.445 | -15.8923
9.077 | -34.4454 | -15.9022
9.0831 | -34.4458 | -15.9121
9.0892 | -34.4462 | -15.922
9.0953 | -34.4466 | -15.9319
9.1015 | -34.4471 | -15.9417
9.1076 | -34.4475 | -15.9513
9.1137 | -34.4479 | -15.9612
9.1198 | -34.4483 | -15.971
9.1259 | -34.4487 | -15.9809
9.132 | -34.4492 | -15.9909
9.1381 | -34.4496 | -16.0005
9.1442 | -34.45 | -16.01
9.1503 | -34.4504 | -16.0199
9.1564 | -34.4509 | -16.0297
9.1625 | -34.4513 | -16.0395
9.1686 | -34.4517 | -16.0492
9.1747 | -34.4521 | -16.059
9.1808 | -34.4526 | -16.0688
9.1869 | -34.453 | -16.0786
9.193 | -34.4534 | -16.0884
9.1991 | -34.4538 | -16.0983
9.2052 | -34.4542 | -16.1081
9.2113 | -34.4547 | -16.118
9.2174 | -34.4551 | -16.1277
9.2235 | -34.4556 | -16.1375
9.2296 | -34.456 | -16.1472
9.2357 | -34.4564 | -16.1571
9.2419 | -34.4568 | -16.1669
9.248 | -34.4573 | -16.1767
9.2541 | -34.4577 | -16.1866
9.2602 | -34.4582 | -16.1965
9.2663 | -34.4586 | -16.206
9.2724 | -34.459 | -16.2157
9.2785 | -34.4594 | -16.2254
9.2846 | -34.4599 | -16.235
9.2907 | -34.4603 | -16.2448
9.2968 | -34.4607 | -16.2548
9.3029 | -34.4612 | -16.2647
9.309 | -34.4616 | -16.2745
9.3151 | -34.462 | -16.2842
9.3212 | -34.4625 | -16.2939
9.3273 | -34.4629 | -16.3036
9.3334 | -34.4633 | -16.3133
9.3395 | -34.4638 | -16.3229
9.3456 | -34.4642 | -16.3325
9.3517 | -34.4646 | -16.3423
9.3578 | -34.465 | -16.3521
9.3639 | -34.4655 | -16.3618
9.37 | -34.4659 | -16.3716
9.3761 | -34.4663 | -16.3815
9.3822 | -34.4668 | -16.3913
9.3884 | -34.4672 | -16.4007
9.3945 | -34.4677 | -16.4105
9.4006 | -34.4681 | -16.4202
9.4067 | -34.4686 | -16.4301
9.4128 | -34.469 | -16.44
9.4189 | -34.4694 | -16.4498
9.425 | -34.4699 | -16.4595
9.4311 | -34.4703 | -16.4693
9.4372 | -34.4708 | -16.479
9.4433 | -34.4712 | -16.4888
9.4494 | -34.4717 | -16.4983
9.4555 | -34.4721 | -16.5082
9.4616 | -34.4726 | -16.5179
9.4677 | -34.473 | -16.5276
9.4738 | -34.4734 | -16.5373
9.4799 | -34.4739 | -16.547
9.486 | -34.4743 | -16.5569
9.4921 | -34.4747 | -16.5665
9.4982 | -34.4752 | -16.5761
9.5043 | -34.4756 | -16.5858
9.5104 | -34.4761 | -16.5954
9.5165 | -34.4765 | -16.6051
9.5226 | -34.4769 | -16.6149
9.5288 | -34.4774 | -16.6244
9.5349 | -34.4778 | -16.6342
9.541 | -34.4783 | -16.6438
9.5471 | -34.4787 | -16.6535
9.5532 | -34.4792 | -16.6632
9.5593 | -34.4796 | -16.673
9.5654 | -34.4801 | -16.6827
9.5715 | -34.4805 | -16.6924
9.5776 | -34.4809 | -16.7022
9.5837 | -34.4813 | -16.712
9.5898 | -34.4818 | -16.7218
9.5959 | -34.4822 | -16.7315
9.602 | -34.4827 | -16.741
9.6081 | -34.4831 | -16.7507
9.6142 | -34.4836 | -16.7603
9.6203 | -34.484 | -16.77
9.6264 | -34.4844 | -16.7797
9.6325 | -34.4848 | -16.7895
9.6386 | -34.4853 | -16.7991
9.6447 | -34.4857 | -16.8088
9.6508 | -34.4862 | -16.8185
9.6569 | -34.4866 | -16.8281
9.663 | -34.4871 | -16.8379
9.6691 | -34.4875 | -16.8476
9.6753 | -34.488 | -16.8573
9.6814 | -34.4884 | -16.8669
9.6875 | -34.4889 | -16.8764
9.6936 | -34.4893 | -16.8859
9.6997 | -34.4898 | -16.8958
9.7058 | -34.4902 | -16.9054
9.7119 | -34.4907 | -16.9152
9.718 | -34.4911 | -16.9251
9.7241 | -34.4916 | -16.9347
9.7302 | -34.492 | -16.9445
9.7363 | -34.4924 | -16.954
9.7424 | -34.4929 | -16.9637
9.7485 | -34.4933 | -16.9734
9.7546 | -34.4938 | -16.9829
9.7607 | -34.4942 | -16.9929
9.7668 | -34.4947 | -17.0026
9.7729 | -34.4951 | -17.0123
9.779 | -34.4956 | -17.0218
9.7851 | -34.4961 | -17.0317
9.7912 | -34.4966 | -17.0413
9.7973 | -34.497 | -17.0512
9.8034 | -34.4975 | -17.0607
9.8095 | -34.4979 | -17.0706
9.8157 | -34.4984 | -17.0804
9.8218 | -34.4989 | -17.0902
9.8279 | -34.4993 | -17.0997
9.834 | -34.4997 | -17.1096
9.8401 | -34.5002 | -17.1192
9.8462 | -34.5006 | -17.1287
9.8523 | -34.5011 | -17.1384
9.8584 | -34.5015 | -17.148
9.8645 | -34.502 | -17.1577
9.8706 | -34.5024 | -17.1674
9.8767 | -34.5028 | -17.1771
9.8828 | -34.5033 | -17.1868
9.8889 | -34.5037 | -17.1962
9.895 | -34.5042 | -17.2059
9.9011 | -34.5047 | -17.2155
9.9072 | -34.5051 | -17.2252
9.9133 | -34.5056 | -17.2349
9.9194 | -34.506 | -17.2445
9.9255 | -34.5065 | -17.254
9.9316 | -34.5069 | -17.2637
9.9377 | -34.5074 | -17.2733
9.9438 | -34.5078 | -17.2831
9.9499 | -34.5083 | -17.293
9.956 | -34.5087 | -17.3026
9.9622 | -34.5092 | -17.3123
9.9683 | -34.5096 | -17.322
9.9744 | -34.51 | -17.3319
9.9805 | -34.5105 | -17.3415
9.9866 | -34.511 | -17.3511
9.9927 | -34.5114 | -17.3606
9.9988 | -34.5118 | -17.3704
10.0049 | -34.5123 | -17.3801
10.011 | -34.5128 | -17.3899
10.0171 | -34.5132 | -17.3994
10.0232 | -34.5137 | -17.4091
10.0293 | -34.5142 | -17.4188
10.0354 | -34.5146 | -17.4285
10.0415 | -34.5151 | -17.4381
10.0476 | -34.5155 | -17.4479
10.0537 | -34.516 | -17.4577
10.0598 | -34.5164 | -17.4674
10.0659 | -34.5169 | -17.4771
10.072 | -34.5173 | -17.4867
10.0781 | -34.5178 | -17.4963
10.0842 | -34.5182 | -17.506
10.0903 | -34.5187 | -17.5157
10.0964 | -34.5192 | -17.5251
10.1026 | -34.5196 | -17.5349
10.1087 | -34.5201 | -17.5445
10.1148 | -34.5205 | -17.5542
10.1209 | -34.5209 | -17.5637
10.127 | -34.5214 | -17.5732
10.1331 | -34.5219 | -17.5827
10.1392 | -34.5223 | -17.5921
10.1453 | -34.5228 | -17.6016
10.1514 | -34.5232 | -17.6114
10.1575 | -34.5237 | -17.6211
10.1636 | -34.5241 | -17.6305
10.1697 | -34.5246 | -17.6403
10.1758 | -34.5251 | -17.6497
10.1819 | -34.5255 | -17.6592
10.188 | -34.5259 | -17.6689
10.1941 | -34.5264 | -17.6786
10.2002 | -34.5268 | -17.6882
10.2063 | -34.5273 | -17.6978
10.2124 | -34.5277 | -17.7075
10.2185 | -34.5282 | -17.7172
10.2246 | -34.5287 | -17.7267
10.2307 | -34.5291 | -17.7364
10.2368 | -34.5296 | -17.7458
10.2429 | -34.53 | -17.7554
10.2491 | -34.5305 | -17.7651
10.2552 | -34.5309 | -17.7748
10.2613 | -34.5314 | -17.7845
10.2674 | -34.5318 | -17.7942
10.2735 | -34.5323 | -17.804
10.2796 | -34.5327 | -17.8136
10.2857 | -34.5332 | -17.8234
10.2918 | -34.5336 | -17.8332
10.2979 | -34.5341 | -17.8428
10.304 | -34.5346 | -17.8521
10.3101 | -34.535 | -17.8618
10.3162 | -34.5355 | -17.8715
10.3223 | -34.5359 | -17.881
10.3284 | -34.5363 | -17.8906
10.3345 | -34.5368 | -17.9
10.3406 | -34.5372 | -17.9095
10.3467 | -34.5377 | -17.9191
10.3528 | -34.5381 | -17.9286
10.3589 | -34.5385 | -17.9382
10.365 | -34.539 | -17.9478
10.3711 | -34.5394 | -17.9575
10.3772 | -34.5399 | -17.967
10.3833 | -34.5404 | -17.9766
10.3895 | -34.5409 | -17.9862
10.3956 | -34.5413 | -17.9958
10.4017 | -34.5418 | -18.0053
10.4078 | -34.5422 | -18.0148
10.4139 | -34.5427 | -18.0245
10.42 | -34.5431 | -18.0342
10.4261 | -34.5436 | -18.0438
10.4322 | -34.544 | -18.0534
10.4383 | -34.5445 | -18.0628
10.4444 | -34.5449 | -18.0722
10.4505 | -34.5453 | -18.0817
10.4566 | -34.5458 | -18.0912
10.4627 | -34.5462 | -18.1007
10.4688 | -34.5467 | -18.1103
10.4749 | -34.5472 | -18.1201
10.481 | -34.5476 | -18.1298
10.4871 | -34.5481 | -18.1393
10.4932 | -34.5486 | -18.1489
10.4993 | -34.549 | -18.1585
10.5054 | -34.5494 | -18.1679
10.5115 | -34.5499 | -18.1776
10.5176 | -34.5504 | -18.187
10.5237 | -34.5508 | -18.1968
10.5298 | -34.5513 | -18.2061
10.536 | -34.5517 | -18.2158
10.5421 | -34.5522 | -18.2253
10.5482 | -34.5526 | -18.2348
10.5543 | -34.5531 | -18.2444
10.5604 | -34.5536 | -18.2538
10.5665 | -34.554 | -18.2633
10.5726 | -34.5544 | -18.2728
10.5787 | -34.5549 | -18.2822
10.5848 | -34.5554 | -18.2913
10.5909 | -34.5559 | -18.301
10.597 | -34.5563 | -18.3106
10.6031 | -34.5568 | -18.3205
10.6092 | -34.5572 | -18.33
10.6153 | -34.5577 | -18.3396
10.6214 | -34.5582 | -18.3491
10.6275 | -34.5586 | -18.3586
10.6336 | -34.5591 | -18.3682
10.6397 | -34.5595 | -18.3778
10.6458 | -34.56 | -18.3872
10.6519 | -34.5604 | -18.3968
10.658 | -34.5609 | -18.4067
10.6641 | -34.5614 | -18.4162
10.6702 | -34.5618 | -18.4256
10.6764 | -34.5623 | -18.4353
10.6825 | -34.5627 | -18.4448
10.6886 | -34.5632 | -18.4541
10.6947 | -34.5637 | -18.4635
10.7008 | -34.5641 | -18.4732
10.7069 | -34.5646 | -18.4828
10.713 | -34.565 | -18.4922
10.7191 | -34.5655 | -18.5017
10.7252 | -34.5659 | -18.5112
10.7313 | -34.5663 | -18.5208
10.7374 | -34.5668 | -18.5302
10.7435 | -34.5672 | -18.5399
10.7496 | -34.5677 | -18.5492
10.7557 | -34.5681 | -18.5589
10.7618 | -34.5686 | -18.5681
10.7679 | -34.5691 | -18.5775
10.774 | -34.5695 | -18.5872
10.7801 | -34.57 | -18.5969
10.7862 | -34.5704 | -18.6064
10.7923 | -34.5709 | -18.6159
10.7984 | -34.5713 | -18.6253
10.8045 | -34.5718 | -18.6347
10.8106 | -34.5723 | -18.6444
10.8168 | -34.5727 | -18.6539
10.8229 | -34.5732 | -18.6635
10.829 | -34.5736 | -18.6729
10.8351 | -34.5741 | -18.6822
10.8412 | -34.5745 | -18.6915
10.8473 | -34.575 | -18.701
10.8534 | -34.5755 | -18.7106
10.8595 | -34.5759 | -18.7201
10.8656 | -34.5764 | -18.7294
10.8717 | -34.5769 | -18.7389
10.8778 | -34.5773 | -18.7484
10.8839 | -34.5778 | -18.7578
10.89 | -34.5783 | -18.7672
10.8961 | -34.5788 | -18.7766
10.9022 | -34.5792 | -18.7861
10.9083 | -34.5797 | -18.7955
10.9144 | -34.5801 | -18.8049
10.9205 | -34.5805 | -18.8143
10.9266 | -34.581 | -18.8237
10.9327 | -34.5815 | -18.8332
10.9388 | -34.582 | -18.8426
10.9449 | -34.5824 | -18.852
10.951 | -34.5829 | -18.8615
10.9571 | -34.5833 | -18.871
10.9633 | -34.5838 | -18.8806
10.9694 | -34.5842 | -18.8903
10.9755 | -34.5847 | -18.8996
10.9816 | -34.5851 | -18.909
10.9877 | -34.5856 | -18.9184
10.9938 | -34.5861 | -18.9279
10.9999 | -34.5866 | -18.9371
11.006 | -34.587 | -18.9466
11.0121 | -34.5875 | -18.9562
11.0182 | -34.588 | -18.9655
11.0243 | -34.5884 | -18.9748
11.0304 | -34.5889 | -18.9841
11.0365 | -34.5893 | -18.9934
11.0426 | -34.5898 | -19.0028
11.0487 | -34.5903 | -19.0124
11.0548 | -34.5908 | -19.0216
11.0609 | -34.5912 | -19.0312
11.067 | -34.5917 | -19.0406
11.0731 | -34.5922 | -19.05
11.0792 | -34.5927 | -19.0594
11.0853 | -34.5932 | -19.0688
11.0914 | -34.5936 | -19.0784
11.0975 | -34.5941 | -19.0877
11.1037 | -34.5945 | -19.0973
11.1098 | -34.5951 | -19.1067
11.1159 | -34.5955 | -19.1162
11.122 | -34.596 | -19.1254
11.1281 | -34.5965 | -19.1349
11.1342 | -34.5969 | -19.1443
11.1403 | -34.5974 | -19.1538
11.1464 | -34.5979 | -19.1633
11.1525 | -34.5983 | -19.1726
11.1586 | -34.5988 | -19.1821
11.1647 | -34.5992 | -19.1915
11.1708 | -34.5997 | -19.2008
11.1769 | -34.6002 | -19.2102
11.183 | -34.6006 | -19.2196
11.1891 | -34.6012 | -19.2291
11.1952 | -34.6017 | -19.2384
11.2013 | -34.6021 | -19.248
11.2074 | -34.6026 | -19.2575
11.2135 | -34.6031 | -19.267
11.2196 | -34.6035 | -19.2764
11.2257 | -34.604 | -19.286
11.2318 | -34.6045 | -19.2953
11.2379 | -34.605 | -19.3049
11.244 | -34.6055 | -19.3145
11.2502 | -34.606 | -19.3239
11.2563 | -34.6065 | -19.3333
11.2624 | -34.607 | -19.3426
11.2685 | -34.6074 | -19.3518
11.2746 | -34.6079 | -19.3611
11.2807 | -34.6084 | -19.3706
11.2868 | -34.6088 | -19.3803
11.2929 | -34.6093 | -19.389
11.299 | -34.6098 | -19.3983
11.3051 | -34.6102 | -19.4076
11.3112 | -34.6107 | -19.4171
11.3173 | -34.6112 | -19.4263
11.3234 | -34.6117 | -19.4355
11.3295 | -34.6121 | -19.4447
11.3356 | -34.6126 | -19.4541
11.3417 | -34.6131 | -19.4634
11.3478 | -34.6135 | -19.4728
11.3539 | -34.614 | -19.4823
11.36 | -34.6145 | -19.4915
11.3661 | -34.6149 | -19.5008
11.3722 | -34.6154 | -19.5101
11.3783 | -34.6159 | -19.5196
11.3844 | -34.6164 | -19.5291
11.3906 | -34.6169 | -19.5383
11.3967 | -34.6174 | -19.5476
11.4028 | -34.6179 | -19.5569
11.4089 | -34.6184 | -19.5662
11.415 | -34.6188 | -19.5754
11.4211 | -34.6192 | -19.5847
11.4272 | -34.6197 | -19.5941
11.4333 | -34.6202 | -19.6032
11.4394 | -34.6207 | -19.6123
11.4455 | -34.6211 | -19.6215
11.4516 | -34.6216 | -19.6307
11.4577 | -34.6221 | -19.6401
11.4638 | -34.6226 | -19.6492
11.4699 | -34.6231 | -19.6586
11.476 | -34.6236 | -19.6681
11.4821 | -34.624 | -19.6773
11.4882 | -34.6246 | -19.6867
11.4943 | -34.625 | -19.6962
11.5004 | -34.6255 | -19.7051
11.5065 | -34.626 | -19.7144
11.5126 | -34.6265 | -19.7236
11.5187 | -34.627 | -19.7328
11.5248 | -34.6274 | -19.7422
11.5309 | -34.6279 | -19.7516
11.5371 | -34.6284 | -19.7607
11.5432 | -34.6289 | -19.7702
11.5493 | -34.6294 | -19.7797
11.5554 | -34.6299 | -19.7891
11.5615 | -34.6303 | -19.7985
11.5676 | -34.6308 | -19.8079
11.5737 | -34.6313 | -19.8169
11.5798 | -34.6318 | -19.826
11.5859 | -34.6323 | -19.8351
11.592 | -34.6328 | -19.8445
11.5981 | -34.6333 | -19.8538
11.6042 | -34.6337 | -19.863
11.6103 | -34.6342 | -19.8721
11.6164 | -34.6347 | -19.8811
11.6225 | -34.6351 | -19.8903
11.6286 | -34.6356 | -19.8995
11.6347 | -34.636 | -19.9087
11.6408 | -34.6366 | -19.9179
11.6469 | -34.637 | -19.9272
11.653 | -34.6375 | -19.9363
11.6591 | -34.638 | -19.9456
11.6652 | -34.6385 | -19.9547
11.6713 | -34.639 | -19.9638
11.6775 | -34.6394 | -19.9728
11.6836 | -34.6399 | -19.9822
11.6897 | -34.6405 | -19.9916
11.6958 | -34.6409 | -20.0008
11.7019 | -34.6414 | -20.0102
11.708 | -34.6419 | -20.0193
11.7141 | -34.6424 | -20.0281
11.7202 | -34.6429 | -20.0375
11.7263 | -34.6434 | -20.0467
11.7324 | -34.6438 | -20.0561
11.7385 | -34.6443 | -20.0655
11.7446 | -34.6448 | -20.0746
11.7507 | -34.6453 | -20.0836
11.7568 | -34.6458 | -20.0928
11.7629 | -34.6462 | -20.1021
11.769 | -34.6467 | -20.1112
11.7751 | -34.6472 | -20.1202
11.7812 | -34.6477 | -20.1292
11.7873 | -34.6482 | -20.1385
11.7934 | -34.6487 | -20.1478
11.7995 | -34.6492 | -20.1569
11.8056 | -34.6497 | -20.166
11.8117 | -34.6502 | -20.1751
11.8178 | -34.6507 | -20.1841
11.824 | -34.6512 | -20.1931
11.8301 | -34.6517 | -20.202
11.8362 | -34.6521 | -20.2112
11.8423 | -34.6526 | -20.2203
11.8484 | -34.6531 | -20.2295
11.8545 | -34.6536 | -20.2384
11.8606 | -34.6541 | -20.2474
11.8667 | -34.6546 | -20.2567
11.8728 | -34.6551 | -20.2658
11.8789 | -34.6555 | -20.275
11.885 | -34.656 | -20.2843
11.8911 | -34.6565 | -20.2931
11.8972 | -34.657 | -20.302
11.9033 | -34.6575 | -20.311
11.9094 | -34.658 | -20.32
11.9155 | -34.6584 | -20.329
11.9216 | -34.659 | -20.3378
11.9277 | -34.6595 | -20.3468
11.9338 | -34.6599 | -20.3559
11.9399 | -34.6604 | -20.3651
11.946 | -34.6609 | -20.3741
11.9521 | -34.6614 | -20.3831
11.9582 | -34.6619 | -20.392
11.9644 | -34.6624 | -20.401
11.9705 | -34.6629 | -20.4101
11.9766 | -34.6634 | -20.4191
11.9827 | -34.6639 | -20.4281
11.9888 | -34.6644 | -20.437
11.9949 | -34.6649 | -20.4457
12.001 | -34.6654 | -20.4548
12.0071 | -34.6659 | -20.4639
12.0132 | -34.6664 | -20.4729
12.0193 | -34.6668 | -20.482
12.0254 | -34.6673 | -20.491
12.0315 | -34.6678 | -20.4998
12.0376 | -34.6683 | -20.5087
12.0437 | -34.6688 | -20.5176
12.0498 | -34.6693 | -20.5267
12.0559 | -34.6698 | -20.5356
12.062 | -34.6703 | -20.5444
12.0681 | -34.6708 | -20.5534
12.0742 | -34.6713 | -20.5622
12.0803 | -34.6718 | -20.5714
12.0864 | -34.6723 | -20.5806
12.0925 | -34.6728 | -20.5896
12.0986 | -34.6733 | -20.5984
12.1047 | -34.6738 | -20.6073
12.1109 | -34.6743 | -20.6161
12.117 | -34.6748 | -20.625
12.1231 | -34.6753 | -20.6338
12.1292 | -34.6758 | -20.6428
12.1353 | -34.6763 | -20.6514
12.1414 | -34.6768 | -20.6605
12.1475 | -34.6772 | -20.6695
12.1536 | -34.6777 | -20.6784
12.1597 | -34.6782 | -20.6873
12.1658 | -34.6787 | -20.6962
12.1719 | -34.6792 | -20.7047
12.178 | -34.6797 | -20.7136
12.1841 | -34.6802 | -20.7226
12.1902 | -34.6807 | -20.7316
12.1963 | -34.6812 | -20.7405
12.2024 | -34.6817 | -20.7491
12.2085 | -34.6822 | -20.758
12.2146 | -34.6827 | -20.767
12.2207 | -34.6832 | -20.7758
12.2268 | -34.6837 | -20.7847
12.2329 | -34.6842 | -20.7935
12.239 | -34.6848 | -20.8019
12.2451 | -34.6853 | -20.8108
12.2513 | -34.6858 | -20.8198
12.2574 | -34.6863 | -20.8286
12.2635 | -34.6868 | -20.8375
12.2696 | -34.6873 | -20.8462
12.2757 | -34.6878 | -20.8548
12.2818 | -34.6883 | -20.8638
12.2879 | -34.6887 | -20.8726
12.294 | -34.6893 | -20.8813
12.3001 | -34.6897 | -20.8901
12.3062 | -34.6903 | -20.8989
12.3123 | -34.6908 | -20.9074
12.3184 | -34.6913 | -20.9161
12.3245 | -34.6918 | -20.9248
12.3306 | -34.6923 | -20.9336
12.3367 | -34.6928 | -20.9424
12.3428 | -34.6933 | -20.9512
12.3489 | -34.6938 | -20.9596
12.355 | -34.6943 | -20.9685
12.3611 | -34.6948 | -20.9774
12.3672 | -34.6952 | -20.9862
12.3733 | -34.6958 | -20.995
12.3794 | -34.6963 | -21.0034
12.3855 | -34.6968 | -21.0121
12.3916 | -34.6973 | -21.0209
12.3978 | -34.6978 | -21.0298
12.4039 | -34.6983 | -21.0386
12.41 | -34.6988 | -21.0474
12.4161 | -34.6993 | -21.0557
12.4222 | -34.6998 | -21.0644
12.4283 | -34.7003 | -21.0731
12.4344 | -34.7008 | -21.0819
12.4405 | -34.7013 | -21.0906
12.4466 | -34.7018 | -21.0992
12.4527 | -34.7023 | -21.1077
12.4588 | -34.7028 | -21.1165
12.4649 | -34.7033 | -21.1253
12.471 | -34.7038 | -21.1341
12.4771 | -34.7042 | -21.1428
12.4832 | -34.7047 | -21.1514
12.4893 | -34.7052 | -21.1599
12.4954 | -34.7057 | -21.1686
12.5015 | -34.7062 | -21.1772
12.5076 | -34.7067 | -21.1861
12.5137 | -34.7073 | -21.1948
12.5198 | -34.7078 | -21.2032
12.5259 | -34.7082 | -21.212
12.532 | -34.7087 | -21.2207
12.5382 | -34.7093 | -21.2295
12.5443 | -34.7097 | -21.2382
12.5504 | -34.7102 | -21.2469
12.5565 | -34.7108 | -21.2553
12.5626 | -34.7113 | -21.264
12.5687 | -34.7117 | -21.2728
12.5748 | -34.7122 | -21.2815
12.5809 | -34.7127 | -21.2902
12.587 | -34.7132 | -21.2989
12.5931 | -34.7138 | -21.3073
12.5992 | -34.7143 | -21.3159
12.6053 | -34.7148 | -21.3245
12.6114 | -34.7152 | -21.3332
12.6175 | -34.7158 | -21.3418
12.6236 | -34.7163 | -21.3505
12.6297 | -34.7168 | -21.359
12.6358 | -34.7173 | -21.3674
12.6419 | -34.7178 | -21.3761
12.648 | -34.7183 | -21.3848
12.6541 | -34.7188 | -21.3934
12.6602 | -34.7193 | -21.4016
12.6663 | -34.7198 | -21.4102
12.6724 | -34.7203 | -21.4188
12.6785 | -34.7208 | -21.4274
12.6847 | -34.7213 | -21.4359
12.6908 | -34.7218 | -21.4445
12.6969 | -34.7223 | -21.4529
12.703 | -34.7228 | -21.4616
12.7091 | -34.7233 | -21.4702
12.7152 | -34.7238 | -21.4786
12.7213 | -34.7242 | -21.4874
12.7274 | -34.7247 | -21.4959
12.7335 | -34.7253 | -21.5043
12.7396 | -34.7258 | -21.5126
12.7457 | -34.7263 | -21.5214
12.7518 | -34.7268 | -21.53
12.7579 | -34.7273 | -21.5386
12.764 | -34.7278 | -21.5472
12.7701 | -34.7282 | -21.5557
12.7762 | -34.7287 | -21.5641
12.7823 | -34.7292 | -21.5726
12.7884 | -34.7297 | -21.581
12.7945 | -34.7302 | -21.5896
12.8006 | -34.7307 | -21.5981
12.8067 | -34.7312 | -21.6066
12.8128 | -34.7317 | -21.6149
12.8189 | -34.7322 | -21.6235
12.8251 | -34.7327 | -21.6322
12.8312 | -34.7332 | -21.6407
12.8373 | -34.7337 | -21.6491
12.8434 | -34.7342 | -21.6576
12.8495 | -34.7347 | -21.6661
12.8556 | -34.7352 | -21.6747
12.8617 | -34.7357 | -21.6831
12.8678 | -34.7361 | -21.6919
12.8739 | -34.7366 | -21.7002
12.88 | -34.7371 | -21.7087
12.8861 | -34.7375 | -21.7172
12.8922 | -34.738 | -21.7258
12.8983 | -34.7385 | -21.7345
12.9044 | -34.739 | -21.7429
12.9105 | -34.7395 | -21.7512
12.9166 | -34.74 | -21.7597
12.9227 | -34.7405 | -21.768
12.9288 | -34.741 | -21.7767
12.9349 | -34.7415 | -21.7852
12.941 | -34.742 | -21.7938
12.9471 | -34.7425 | -21.8022
12.9532 | -34.743 | -21.8107
12.9593 | -34.7435 | -21.8193
12.9654 | -34.744 | -21.8278
12.9716 | -34.7445 | -21.8364
12.9777 | -34.745 | -21.8449
12.9838 | -34.7455 | -21.8533
12.9899 | -34.746 | -21.8619
12.996 | -34.7464 | -21.8704
13.0021 | -34.747 | -21.879
13.0082 | -34.7475 | -21.8873
13.0143 | -34.748 | -21.8954
13.0204 | -34.7485 | -21.9037
13.0265 | -34.749 | -21.9121
13.0326 | -34.7495 | -21.9206
13.0387 | -34.75 | -21.9291
13.0448 | -34.7505 | -21.9376
13.0509 | -34.7509 | -21.946
13.057 | -34.7515 | -21.9544
13.0631 | -34.752 | -21.9629
13.0692 | -34.7525 | -21.9713
13.0753 | -34.753 | -21.98
13.0814 | -34.7535 | -21.9886
13.0875 | -34.754 | -21.9969
13.0936 | -34.7545 | -22.0053
13.0997 | -34.755 | -22.0138
13.1058 | -34.7555 | -22.0223
13.112 | -34.756 | -22.0308
13.1181 | -34.7565 | -22.0392
13.1242 | -34.757 | -22.0477
13.1303 | -34.7575 | -22.0561
13.1364 | -34.758 | -22.0646
13.1425 | -34.7585 | -22.0729
13.1486 | -34.759 | -22.0813
13.1547 | -34.7594 | -22.0898
13.1608 | -34.7599 | -22.0983
13.1669 | -34.7604 | -22.1067
13.173 | -34.761 | -22.1152
13.1791 | -34.7614 | -22.1236
13.1852 | -34.7619 | -22.132
13.1913 | -34.7624 | -22.14
13.1974 | -34.7629 | -22.1487
13.2035 | -34.7633 | -22.1571
13.2096 | -34.7638 | -22.1653
13.2157 | -34.7643 | -22.1737
13.2218 | -34.7648 | -22.1821
13.2279 | -34.7653 | -22.1903
13.234 | -34.7658 | -22.1988
13.2401 | -34.7664 | -22.2071
13.2462 | -34.7668 | -22.2155
13.2524 | -34.7673 | -22.2239
13.2585 | -34.7678 | -22.2324
13.2646 | -34.7683 | -22.2408
13.2707 | -34.7688 | -22.2492
13.2768 | -34.7693 | -22.2577
13.2829 | -34.7698 | -22.2661
13.289 | -34.7703 | -22.2744
13.2951 | -34.7708 | -22.2826
13.3012 | -34.7713 | -22.2912
13.3073 | -34.7718 | -22.2996
13.3134 | -34.7723 | -22.3081
13.3195 | -34.7727 | -22.3163
13.3256 | -34.7732 | -22.3246
13.3317 | -34.7737 | -22.333
13.3378 | -34.7742 | -22.3415
13.3439 | -34.7747 | -22.3501
13.35 | -34.7752 | -22.3589
13.3561 | -34.7756 | -22.3671
13.3622 | -34.7761 | -22.3755
13.3683 | -34.7765 | -22.3839
13.3744 | -34.777 | -22.3922
13.3805 | -34.7775 | -22.4006
13.3866 | -34.778 | -22.4088
13.3927 | -34.7784 | -22.4172
13.3989 | -34.7789 | -22.4258
13.405 | -34.7794 | -22.4343
13.4111 | -34.7799 | -22.4426
13.4172 | -34.7803 | -22.4511
13.4233 | -34.7808 | -22.4596
13.4294 | -34.7813 | -22.468
13.4355 | -34.7818 | -22.4762
13.4416 | -34.7824 | -22.4845
13.4477 | -34.7829 | -22.4928
13.4538 | -34.7834 | -22.5013
13.4599 | -34.7839 | -22.5097
13.466 | -34.7844 | -22.5182
13.4721 | -34.7848 | -22.5265
13.4782 | -34.7853 | -22.535
13.4843 | -34.7858 | -22.5434
13.4904 | -34.7864 | -22.552
13.4965 | -34.7868 | -22.5602
13.5026 | -34.7873 | -22.5684
13.5087 | -34.7878 | -22.5768
13.5148 | -34.7883 | -22.5852
13.5209 | -34.7887 | -22.5938
13.527 | -34.7892 | -22.6024
13.5331 | -34.7898 | -22.6109
13.5393 | -34.7902 | -22.6191
13.5454 | -34.7907 | -22.6274
13.5515 | -34.7912 | -22.6359
13.5576 | -34.7918 | -22.6444
13.5637 | -34.7923 | -22.6533
13.5698 | -34.7927 | -22.6618
13.5759 | -34.7932 | -22.67
13.582 | -34.7937 | -22.6785
13.5881 | -34.7941 | -22.6868
13.5942 | -34.7946 | -22.6953
13.6003 | -34.7951 | -22.7036
13.6064 | -34.7956 | -22.7122
13.6125 | -34.7961 | -22.7206
13.6186 | -34.7966 | -22.7292
13.6247 | -34.7971 | -22.7375
13.6308 | -34.7976 | -22.7462
13.6369 | -34.7981 | -22.7548
13.643 | -34.7985 | -22.7633
13.6491 | -34.799 | -22.7717
13.6552 | -34.7995 | -22.7803
13.6613 | -34.8 | -22.7886
13.6674 | -34.8005 | -22.7971
13.6735 | -34.801 | -22.8055
13.6796 | -34.8015 | -22.8138
13.6858 | -34.802 | -22.8223
13.6919 | -34.8025 | -22.8309
13.698 | -34.803 | -22.8394
13.7041 | -34.8035 | -22.848
13.7102 | -34.804 | -22.8564
13.7163 | -34.8045 | -22.8648
13.7224 | -34.805 | -22.8733
13.7285 | -34.8055 | -22.8817
13.7346 | -34.8059 | -22.8902
13.7407 | -34.8064 | -22.8986
13.7468 | -34.8069 | -22.9072
13.7529 | -34.8074 | -22.9157
13.759 | -34.8079 | -22.9242
13.7651 | -34.8084 | -22.9327
13.7712 | -34.8089 | -22.9412
13.7773 | -34.8094 | -22.9497
13.7834 | -34.8099 | -22.9581
13.7895 | -34.8104 | -22.9667
13.7956 | -34.8108 | -22.9753
13.8017 | -34.8113 | -22.9838
13.8078 | -34.8118 | -22.9921
13.8139 | -34.8123 | -23.0005
13.82 | -34.8128 | -23.0089
13.8262 | -34.8132 | -23.0175
13.8323 | -34.8138 | -23.026
13.8384 | -34.8142 | -23.0344
13.8445 | -34.8147 | -23.0429
13.8506 | -34.8152 | -23.0514
13.8567 | -34.8157 | -23.0598
13.8628 | -34.8162 | -23.0683
13.8689 | -34.8167 | -23.0769
13.875 | -34.8171 | -23.0854
13.8811 | -34.8176 | -23.0939
13.8872 | -34.8181 | -23.1022
13.8933 | -34.8186 | -23.1105
13.8994 | -34.819 | -23.119
13.9055 | -34.8195 | -23.1276
13.9116 | -34.82 | -23.1359
13.9177 | -34.8205 | -23.1443
13.9238 | -34.821 | -23.1526
13.9299 | -34.8215 | -23.1611
13.936 | -34.8219 | -23.1695
13.9421 | -34.8224 | -23.178
13.9482 | -34.8229 | -23.1865
13.9543 | -34.8233 | -23.195
13.9604 | -34.8238 | -23.2035
13.9665 | -34.8243 | -23.2119
13.9727 | -34.8248 | -23.2205
13.9788 | -34.8253 | -23.2289
13.9849 | -34.8258 | -23.2375
13.991 | -34.8263 | -23.246
13.9971 | -34.8268 | -23.2544
14.0032 | -34.8272 | -23.2628
14.0093 | -34.8277 | -23.2713
14.0154 | -34.8282 | -23.2798
14.0215 | -34.8286 | -23.2882
14.0276 | -34.8291 | -23.2967
14.0337 | -34.8297 | -23.3053
14.0398 | -34.8302 | -23.3137
14.0459 | -34.8306 | -23.3221
14.052 | -34.8311 | -23.3307
14.0581 | -34.8316 | -23.3392
14.0642 | -34.832 | -23.3477
14.0703 | -34.8326 | -23.3563
14.0764 | -34.833 | -23.3647
14.0825 | -34.8335 | -23.3733
14.0886 | -34.834 | -23.3818
14.0947 | -34.8345 | -23.3903
14.1008 | -34.835 | -23.3987
14.1069 | -34.8355 | -23.4072
14.1131 | -34.8359 | -23.4159
14.1192 | -34.8364 | -23.4243
14.1253 | -34.8369 | -23.4327
14.1314 | -34.8374 | -23.4412
14.1375 | -34.8379 | -23.4496
14.1436 | -34.8384 | -23.4581
14.1497 | -34.8389 | -23.4665
14.1558 | -34.8394 | -23.4751
14.1619 | -34.8399 | -23.4836
14.168 | -34.8404 | -23.4919
14.1741 | -34.8409 | -23.5004
14.1802 | -34.8414 | -23.5088
14.1863 | -34.8419 | -23.5173
14.1924 | -34.8424 | -23.5258
14.1985 | -34.8429 | -23.5343
14.2046 | -34.8434 | -23.543
14.2107 | -34.8439 | -23.5514
14.2168 | -34.8443 | -23.56
14.2229 | -34.8448 | -23.5685
14.229 | -34.8453 | -23.5769
14.2351 | -34.8458 | -23.5855
14.2412 | -34.8463 | -23.5939
14.2473 | -34.8468 | -23.6024
14.2534 | -34.8473 | -23.6107
14.2596 | -34.8477 | -23.6192
14.2657 | -34.8482 | -23.6277
14.2718 | -34.8487 | -23.6362
14.2779 | -34.8492 | -23.6447
14.284 | -34.8497 | -23.6534
14.2901 | -34.8502 | -23.6619
14.2962 | -34.8507 | -23.6703
14.3023 | -34.8512 | -23.6788
14.3084 | -34.8516 | -23.6873
14.3145 | -34.8521 | -23.6958
14.3206 | -34.8526 | -23.7045
14.3267 | -34.8531 | -23.7129
14.3328 | -34.8536 | -23.7213
14.3389 | -34.8541 | -23.7297
14.345 | -34.8546 | -23.7383
14.3511 | -34.8551 | -23.7468
14.3572 | -34.8556 | -23.7554
14.3633 | -34.8561 | -23.7639
14.3694 | -34.8566 | -23.7724
14.3755 | -34.8571 | -23.781
14.3816 | -34.8576 | -23.7897
14.3877 | -34.8581 | -23.798
14.3938 | -34.8586 | -23.8068
14.4 | -34.8591 | -23.8154
14.4061 | -34.8596 | -23.8239
14.4122 | -34.8601 | -23.8324
14.4183 | -34.8606 | -23.8409
14.4244 | -34.8611 | -23.8495
14.4305 | -34.8616 | -23.858
14.4366 | -34.8621 | -23.8665
14.4427 | -34.8626 | -23.8752
14.4488 | -34.8631 | -23.8838
14.4549 | -34.8636 | -23.8924
14.461 | -34.8641 | -23.9008
14.4671 | -34.8646 | -23.9094
14.4732 | -34.8651 | -23.918
14.4793 | -34.8656 | -23.9264
14.4854 | -34.8661 | -23.9349
14.4915 | -34.8666 | -23.9432
14.4976 | -34.8671 | -23.9516
14.5037 | -34.8676 | -23.96
14.5098 | -34.8681 | -23.9684
14.5159 | -34.8686 | -23.9769
14.522 | -34.8691 | -23.9853
14.5281 | -34.8696 | -23.9936
14.5342 | -34.8701 | -24.0022
14.5403 | -34.8706 | -24.0106
14.5465 | -34.8711 | -24.019
14.5526 | -34.8716 | -24.0277
14.5587 | -34.8721 | -24.0362
14.5648 | -34.8726 | -24.0444
14.5709 | -34.8731 | -24.0527
14.577 | -34.8736 | -24.0612
14.5831 | -34.8741 | -24.0695
14.5892 | -34.8746 | -24.0781
14.5953 | -34.8751 | -24.0863
14.6014 | -34.8756 | -24.0949
14.6075 | -34.8761 | -24.1034
14.6136 | -34.8766 | -24.112
14.6197 | -34.8771 | -24.1204
14.6258 | -34.8776 | -24.129
14.6319 | -34.8781 | -24.1373
14.638 | -34.8786 | -24.1456
14.6441 | -34.8791 | -24.1543
14.6502 | -34.8796 | -24.1628
14.6563 | -34.8801 | -24.1714
14.6624 | -34.8806 | -24.1799
14.6685 | -34.8811 | -24.1884
14.6746 | -34.8816 | -24.1969
14.6807 | -34.8821 | -24.2054
14.6869 | -34.8826 | -24.2139
14.693 | -34.8832 | -24.2224
14.6991 | -34.8837 | -24.2309
14.7052 | -34.8842 | -24.2393
14.7113 | -34.8847 | -24.2479
14.7174 | -34.8852 | -24.2565
14.7235 | -34.8857 | -24.2648
14.7296 | -34.8862 | -24.2734
14.7357 | -34.8867 | -24.2818
14.7418 | -34.8873 | -24.2902
14.7479 | -34.8878 | -24.2986
14.754 | -34.8883 | -24.3071
14.7601 | -34.8888 | -24.3154
14.7662 | -34.8893 | -24.3239
14.7723 | -34.8898 | -24.3319
14.7784 | -34.8903 | -24.3405
14.7845 | -34.8908 | -24.349
14.7906 | -34.8913 | -24.3576
14.7967 | -34.8918 | -24.3659
14.8028 | -34.8923 | -24.3743
14.8089 | -34.8928 | -24.3826
14.815 | -34.8934 | -24.3911
14.8211 | -34.8938 | -24.3995
14.8272 | -34.8943 | -24.408
14.8334 | -34.8948 | -24.4166
14.8395 | -34.8954 | -24.425
14.8456 | -34.8959 | -24.4335
14.8517 | -34.8964 | -24.442
14.8578 | -34.8969 | -24.4506
14.8639 | -34.8974 | -24.4591
14.87 | -34.8979 | -24.4675
14.8761 | -34.8985 | -24.4758
14.8822 | -34.899 | -24.4842
14.8883 | -34.8994 | -24.4926
14.8944 | -34.8999 | -24.5007
14.9005 | -34.9004 | -24.5092
14.9066 | -34.9009 | -24.5177
14.9127 | -34.9015 | -24.526
14.9188 | -34.902 | -24.5344
14.9249 | -34.9025 | -24.5428
14.931 | -34.903 | -24.5516
14.9371 | -34.9036 | -24.5598
14.9432 | -34.9041 | -24.5682
14.9493 | -34.9046 | -24.5764
14.9554 | -34.9051 | -24.5849
14.9615 | -34.9056 | -24.5933
14.9676 | -34.9061 | -24.602
14.9738 | -34.9066 | -24.6103
14.9799 | -34.9071 | -24.6189
14.986 | -34.9076 | -24.6272
14.9921 | -34.9081 | -24.6356
14.9982 | -34.9086 | -24.6439
15.0043 | -34.9091 | -24.6521
15.0104 | -34.9096 | -24.6606
15.0165 | -34.9101 | -24.6691
15.0226 | -34.9107 | -24.6773
15.0287 | -34.9112 | -24.6858
15.0348 | -34.9117 | -24.6944
15.0409 | -34.9122 | -24.7029
15.047 | -34.9127 | -24.7114
15.0531 | -34.9132 | -24.7196
15.0592 | -34.9137 | -24.7281
15.0653 | -34.9143 | -24.7365
15.0714 | -34.9148 | -24.7447
15.0775 | -34.9153 | -24.753
15.0836 | -34.9158 | -24.7615
15.0897 | -34.9164 | -24.7697
15.0958 | -34.9169 | -24.778
15.1019 | -34.9174 | -24.7864
15.108 | -34.9179 | -24.7946
15.1141 | -34.9185 | -24.8029
15.1203 | -34.919 | -24.8115
15.1264 | -34.9195 | -24.8198
15.1325 | -34.92 | -24.8282
15.1386 | -34.9206 | -24.8365
15.1447 | -34.9211 | -24.8448
15.1508 | -34.9216 | -24.853
15.1569 | -34.9221 | -24.8617
15.163 | -34.9227 | -24.8697
15.1691 | -34.9232 | -24.8781
15.1752 | -34.9237 | -24.8864
15.1813 | -34.9242 | -24.8947
15.1874 | -34.9248 | -24.9033
15.1935 | -34.9253 | -24.9114
15.1996 | -34.9258 | -24.9196
15.2057 | -34.9263 | -24.928
15.2118 | -34.9269 | -24.9362
15.2179 | -34.9274 | -24.9446
15.224 | -34.9279 | -24.9529
15.2301 | -34.9285 | -24.961
15.2362 | -34.929 | -24.9695
15.2423 | -34.9295 | -24.9779
15.2484 | -34.93 | -24.9862
15.2545 | -34.9306 | -24.9946
15.2607 | -34.9311 | -25.0031
15.2668 | -34.9317 | -25.0112
15.2729 | -34.9323 | -25.0194
15.279 | -34.9328 | -25.0279
15.2851 | -34.9333 | -25.0361
15.2912 | -34.9338 | -25.0445
15.2973 | -34.9344 | -25.0529
15.3034 | -34.9349 | -25.0611
15.3095 | -34.9354 | -25.0695
15.3156 | -34.936 | -25.0777
15.3217 | -34.9365 | -25.0858
15.3278 | -34.937 | -25.0941
15.3339 | -34.9376 | -25.1021
15.34 | -34.9381 | -25.1104
15.3461 | -34.9386 | -25.1186
15.3522 | -34.9391 | -25.1269
15.3583 | -34.9397 | -25.1353
15.3644 | -34.9402 | -25.1433
15.3705 | -34.9408 | -25.1513
15.3766 | -34.9413 | -25.1597
15.3827 | -34.9418 | -25.1678
15.3888 | -34.9424 | -25.1762
15.3949 | -34.9429 | -25.1844
15.401 | -34.9434 | -25.1928
15.4072 | -34.944 | -25.2009
15.4133 | -34.9445 | -25.2093
15.4194 | -34.945 | -25.2175
15.4255 | -34.9455 | -25.2256
15.4316 | -34.946 | -25.2339
15.4377 | -34.9466 | -25.2423
15.4438 | -34.9471 | -25.2505
15.4499 | -34.9477 | -25.2588
15.456 | -34.9482 | -25.2671
15.4621 | -34.9487 | -25.2753
15.4682 | -34.9493 | -25.2838
15.4743 | -34.9498 | -25.2917
15.4804 | -34.9503 | -25.3
15.4865 | -34.9508 | -25.3083
15.4926 | -34.9514 | -25.3166
15.4987 | -34.9519 | -25.3248
15.5048 | -34.9524 | -25.3331
15.5109 | -34.953 | -25.3411
15.517 | -34.9535 | -25.3494
15.5231 | -34.9541 | -25.3578
15.5292 | -34.9546 | -25.366
15.5353 | -34.9551 | -25.3743
15.5414 | -34.9557 | -25.3824
15.5476 | -34.9562 | -25.3903
15.5537 | -34.9567 | -25.3986
15.5598 | -34.9573 | -25.4071
15.5659 | -34.9578 | -25.4152
15.572 | -34.9584 | -25.4234
15.5781 | -34.9589 | -25.4315
15.5842 | -34.9594 | -25.4396
15.5903 | -34.96 | -25.4477
15.5964 | -34.9605 | -25.456
15.6025 | -34.961 | -25.4642
15.6086 | -34.9616 | -25.4726
15.6147 | -34.9621 | -25.4809
15.6208 | -34.9627 | -25.4888
15.6269 | -34.9632 | -25.4971
15.633 | -34.9638 | -25.5054
15.6391 | -34.9643 | -25.5135
15.6452 | -34.9648 | -25.5217
15.6513 | -34.9654 | -25.5296
15.6574 | -34.966 | -25.5379
15.6635 | -34.9665 | -25.5462
15.6696 | -34.9671 | -25.5546
15.6757 | -34.9676 | -25.5629
15.6818 | -34.9681 | -25.5712
15.688 | -34.9687 | -25.5793
15.6941 | -34.9692 | -25.5874
15.7002 | -34.9697 | -25.5955
15.7063 | -34.9702 | -25.6035
15.7124 | -34.9708 | -25.612
15.7185 | -34.9713 | -25.6201
15.7246 | -34.9719 | -25.6281
15.7307 | -34.9724 | -25.6363
15.7368 | -34.9729 | -25.6445
15.7429 | -34.9735 | -25.6526
15.749 | -34.9741 | -25.6608
15.7551 | -34.9746 | -25.669
15.7612 | -34.9751 | -25.6769
15.7673 | -34.9757 | -25.6851
15.7734 | -34.9762 | -25.6932
15.7795 | -34.9768 | -25.7013
15.7856 | -34.9773 | -25.7094
15.7917 | -34.9778 | -25.7172
15.7978 | -34.9784 | -25.7253
15.8039 | -34.9789 | -25.7334
15.81 | -34.9794 | -25.7416
15.8161 | -34.98 | -25.7497
15.8222 | -34.9806 | -25.7577
15.8283 | -34.9811 | -25.7655
15.8345 | -34.9817 | -25.7737
15.8406 | -34.9822 | -25.7818
15.8467 | -34.9828 | -25.7899
15.8528 | -34.9833 | -25.798
15.8589 | -34.9839 | -25.8061
15.865 | -34.9845 | -25.814
15.8711 | -34.985 | -25.8222
15.8772 | -34.9855 | -25.8303
15.8833 | -34.9861 | -25.8384
15.8894 | -34.9866 | -25.8466
15.8955 | -34.9872 | -25.8547
15.9016 | -34.9877 | -25.8625
15.9077 | -34.9882 | -25.8706
15.9138 | -34.9887 | -25.8789
15.9199 | -34.9893 | -25.8869
15.926 | -34.9899 | -25.8948
15.9321 | -34.9904 | -25.9027
15.9382 | -34.9909 | -25.9107
15.9443 | -34.9915 | -25.9189
15.9504 | -34.992 | -25.9269
15.9565 | -34.9926 | -25.935
15.9626 | -34.9931 | -25.9431
15.9687 | -34.9936 | -25.9509
15.9749 | -34.9942 | -25.9588
15.981 | -34.9947 | -25.9668
15.9871 | -34.9953 | -25.9749
15.9932 | -34.9958 | -25.9829
15.9993 | -34.9964 | -25.991
16.0054 | -34.997 | -25.9989
16.0115 | -34.9975 | -26.0071
16.0176 | -34.9981 | -26.0152
16.0237 | -34.9986 | -26.0234
16.0298 | -34.9991 | -26.0314
16.0359 | -34.9997 | -26.0395
16.042 | -35.0003 | -26.0474
16.0481 | -35.0008 | -26.0558
16.0542 | -35.0014 | -26.064
16.0603 | -35.0019 | -26.072
16.0664 | -35.0024 | -26.0803
16.0725 | -35.003 | -26.0881
16.0786 | -35.0036 | -26.096
16.0847 | -35.0041 | -26.1041
16.0908 | -35.0046 | -26.1122
16.0969 | -35.0051 | -26.1204
16.103 | -35.0057 | -26.1286
16.1091 | -35.0063 | -26.1364
16.1152 | -35.0068 | -26.1445
16.1214 | -35.0074 | -26.1526
16.1275 | -35.0079 | -26.1607
16.1336 | -35.0085 | -26.1686
16.1397 | -35.009 | -26.1767
16.1458 | -35.0096 | -26.1844
16.1519 | -35.0101 | -26.1926
16.158 | -35.0107 | -26.2006
16.1641 | -35.0112 | -26.2086
16.1702 | -35.0118 | -26.2166
16.1763 | -35.0123 | -26.2246
16.1824 | -35.0129 | -26.2324
16.1885 | -35.0134 | -26.2404
16.1946 | -35.014 | -26.2484
16.2007 | -35.0145 | -26.2563
16.2068 | -35.0151 | -26.2644
16.2129 | -35.0156 | -26.2725
16.219 | -35.0162 | -26.2801
16.2251 | -35.0167 | -26.2882
16.2312 | -35.0173 | -26.2963
16.2373 | -35.0178 | -26.3041
16.2434 | -35.0184 | -26.3121
16.2495 | -35.019 | -26.32
16.2556 | -35.0195 | -26.3281
16.2618 | -35.0201 | -26.3362
16.2679 | -35.0206 | -26.3443
16.274 | -35.0211 | -26.3523
16.2801 | -35.0216 | -26.3603
16.2862 | -35.0222 | -26.3681
16.2923 | -35.0228 | -26.3761
16.2984 | -35.0233 | -26.3841
16.3045 | -35.0238 | -26.3921
16.3106 | -35.0244 | -26.4001
16.3167 | -35.0249 | -26.4081
16.3228 | -35.0255 | -26.4158
16.3289 | -35.026 | -26.4238
16.335 | -35.0266 | -26.4316
16.3411 | -35.0271 | -26.4396
16.3472 | -35.0277 | -26.4476
16.3533 | -35.0282 | -26.4556
16.3594 | -35.0287 | -26.4633
16.3655 | -35.0293 | -26.4711
16.3716 | -35.0298 | -26.479
16.3777 | -35.0304 | -26.487
16.3838 | -35.031 | -26.4948
16.3899 | -35.0315 | -26.5026
16.396 | -35.0321 | -26.5107
16.4021 | -35.0327 | -26.5186
16.4083 | -35.0332 | -26.5263
16.4144 | -35.0337 | -26.5342
16.4205 | -35.0343 | -26.5422
16.4266 | -35.0348 | -26.5498
16.4327 | -35.0354 | -26.5576
16.4388 | -35.0359 | -26.5655
16.4449 | -35.0365 | -26.5735
16.451 | -35.037 | -26.5815
16.4571 | -35.0375 | -26.5895
16.4632 | -35.0381 | -26.597
16.4693 | -35.0386 | -26.605
16.4754 | -35.0392 | -26.6129
16.4815 | -35.0397 | -26.6207
16.4876 | -35.0403 | -26.6287
16.4937 | -35.0408 | -26.6364
16.4998 | -35.0414 | -26.6442
16.5059 | -35.042 | -26.6521
16.512 | -35.0425 | -26.6604
16.5181 | -35.0431 | -26.6683
16.5242 | -35.0437 | -26.6763
16.5303 | -35.0443 | -26.6838
16.5364 | -35.0448 | -26.6917
16.5425 | -35.0454 | -26.6994
16.5487 | -35.0459 | -26.7071
16.5548 | -35.0464 | -26.7149
16.5609 | -35.047 | -26.7228
16.567 | -35.0475 | -26.7305
16.5731 | -35.0481 | -26.7384
16.5792 | -35.0486 | -26.7463
16.5853 | -35.0492 | -26.7543
16.5914 | -35.0497 | -26.7621
16.5975 | -35.0503 | -26.7701
16.6036 | -35.0508 | -26.7777
16.6097 | -35.0513 | -26.7856
16.6158 | -35.0519 | -26.7932
16.6219 | -35.0525 | -26.8012
16.628 | -35.053 | -26.809
16.6341 | -35.0535 | -26.8167
16.6402 | -35.0541 | -26.8244
16.6463 | -35.0546 | -26.8324
16.6524 | -35.0551 | -26.8402
16.6585 | -35.0557 | -26.8479
16.6646 | -35.0563 | -26.8557
16.6707 | -35.0568 | -26.8636
16.6768 | -35.0573 | -26.8714
16.6829 | -35.0579 | -26.879
16.689 | -35.0585 | -26.8869
16.6952 | -35.059 | -26.8946
16.7013 | -35.0595 | -26.9024
16.7074 | -35.0601 | -26.9099
16.7135 | -35.0606 | -26.9178
16.7196 | -35.0612 | -26.9257
16.7257 | -35.0617 | -26.9334
16.7318 | -35.0623 | -26.9413
16.7379 | -35.0628 | -26.9489
16.744 | -35.0634 | -26.9566
16.7501 | -35.0639 | -26.9644
16.7562 | -35.0645 | -26.9722
16.7623 | -35.065 | -26.9798
16.7684 | -35.0655 | -26.9877
16.7745 | -35.0661 | -26.9957
16.7806 | -35.0667 | -27.0031
16.7867 | -35.0672 | -27.011
16.7928 | -35.0678 | -27.0187
16.7989 | -35.0683 | -27.0265
16.805 | -35.0689 | -27.0344
16.8111 | -35.0694 | -27.042
16.8172 | -35.0699 | -27.0501
16.8233 | -35.0705 | -27.0578
16.8294 | -35.0711 | -27.0657
16.8356 | -35.0716 | -27.0735
16.8417 | -35.0722 | -27.0812
16.8478 | -35.0728 | -27.0888
16.8539 | -35.0733 | -27.0964
16.86 | -35.0738 | -27.1043
16.8661 | -35.0744 | -27.112
16.8722 | -35.0749 | -27.1199
16.8783 | -35.0755 | -27.1277
16.8844 | -35.076 | -27.1353
16.8905 | -35.0766 | -27.1429
16.8966 | -35.0772 | -27.1508
16.9027 | -35.0777 | -27.1587
16.9088 | -35.0783 | -27.1665
16.9149 | -35.0788 | -27.1745
16.921 | -35.0794 | -27.1819
16.9271 | -35.0799 | -27.1899
16.9332 | -35.0805 | -27.1977
16.9393 | -35.081 | -27.2056
16.9454 | -35.0815 | -27.2133
16.9515 | -35.0821 | -27.2211
16.9576 | -35.0826 | -27.2285
16.9637 | -35.0832 | -27.2362
16.9698 | -35.0837 | -27.2439
16.9759 | -35.0843 | -27.2516
16.9821 | -35.0849 | -27.2596
16.9882 | -35.0854 | -27.2671
16.9943 | -35.086 | -27.2748
17.0004 | -35.0865 | -27.2826
17.0065 | -35.0871 | -27.2903
17.0126 | -35.0876 | -27.298
17.0187 | -35.0882 | -27.3059
17.0248 | -35.0887 | -27.3135
17.0309 | -35.0892 | -27.3212
17.037 | -35.0898 | -27.329
17.0431 | -35.0903 | -27.3368
17.0492 | -35.0909 | -27.3445
17.0553 | -35.0914 | -27.352
17.0614 | -35.092 | -27.3599
17.0675 | -35.0925 | -27.3675
17.0736 | -35.0931 | -27.3753
17.0797 | -35.0936 | -27.383
17.0858 | -35.0942 | -27.3907
17.0919 | -35.0947 | -27.3985
17.098 | -35.0953 | -27.406
17.1041 | -35.0959 | -27.4135
17.1102 | -35.0964 | -27.4212
17.1163 | -35.097 | -27.429
17.1225 | -35.0975 | -27.4366
17.1286 | -35.0981 | -27.4441
17.1347 | -35.0987 | -27.4518
17.1408 | -35.0992 | -27.4596
17.1469 | -35.0997 | -27.4673
17.153 | -35.1003 | -27.4749
17.1591 | -35.1008 | -27.4827
17.1652 | -35.1014 | -27.4901
17.1713 | -35.1019 | -27.4977
17.1774 | -35.1025 | -27.5053
17.1835 | -35.103 | -27.513
17.1896 | -35.1035 | -27.5206
17.1957 | -35.1041 | -27.5283
17.2018 | -35.1047 | -27.5357
17.2079 | -35.1052 | -27.5434
17.214 | -35.1058 | -27.5512
17.2201 | -35.1063 | -27.5589
17.2262 | -35.1069 | -27.5668
17.2323 | -35.1074 | -27.5744
17.2384 | -35.108 | -27.582
17.2445 | -35.1086 | -27.5897
17.2506 | -35.1091 | -27.5974
17.2567 | -35.1096 | -27.6051
17.2628 | -35.1102 | -27.6127
17.269 | -35.1107 | -27.6201
17.2751 | -35.1113 | -27.6278
17.2812 | -35.1118 | -27.6357
17.2873 | -35.1124 | -27.6433
17.2934 | -35.1129 | -27.6511
17.2995 | -35.1135 | -27.6589
17.3056 | -35.1141 | -27.6663
17.3117 | -35.1146 | -27.6741
17.3178 | -35.1151 | -27.6817
17.3239 | -35.1157 | -27.6894
17.33 | -35.1162 | -27.6971
17.3361 | -35.1168 | -27.7047
17.3422 | -35.1173 | -27.7122
17.3483 | -35.1179 | -27.7199
17.3544 | -35.1184 | -27.7276
17.3605 | -35.1189 | -27.7352
17.3666 | -35.1195 | -27.7428
17.3727 | -35.12 | -27.7505
17.3788 | -35.1206 | -27.7579
17.3849 | -35.1212 | -27.7654
17.391 | -35.1217 | -27.7732
17.3971 | -35.1222 | -27.7808
17.4032 | -35.1227 | -27.7884
17.4094 | -35.1233 | -27.7959
17.4155 | -35.1238 | -27.8035
17.4216 | -35.1244 | -27.8112
17.4277 | -35.1249 | -27.8188
17.4338 | -35.1255 | -27.8265
17.4399 | -35.126 | -27.8342
17.446 | -35.1266 | -27.8415
17.4521 | -35.1272 | -27.849
17.4582 | -35.1277 | -27.8565
17.4643 | -35.1283 | -27.8641
17.4704 | -35.1288 | -27.8717
17.4765 | -35.1294 | -27.8794
17.4826 | -35.1299 | -27.8869
17.4887 | -35.1305 | -27.8946
17.4948 | -35.131 | -27.9022
17.5009 | -35.1315 | -27.9099
17.507 | -35.1321 | -27.9174
17.5131 | -35.1327 | -27.9251
17.5192 | -35.1332 | -27.9324
17.5253 | -35.1338 | -27.94
17.5314 | -35.1343 | -27.9476
17.5375 | -35.1348 | -27.9553
17.5436 | -35.1353 | -27.963
17.5497 | -35.1359 | -27.9706
17.5559 | -35.1365 | -27.9779
17.562 | -35.137 | -27.9854
17.5681 | -35.1376 | -27.993
17.5742 | -35.1381 | -28.0008
17.5803 | -35.1387 | -28.0083
17.5864 | -35.1392 | -28.0158
17.5925 | -35.1398 | -28.0235
17.5986 | -35.1403 | -28.0308
17.6047 | -35.1409 | -28.0385
17.6108 | -35.1414 | -28.046
17.6169 | -35.142 | -28.0536
17.623 | -35.1425 | -28.061
17.6291 | -35.1431 | -28.0685
17.6352 | -35.1436 | -28.0762
17.6413 | -35.1442 | -28.0838
17.6474 | -35.1448 | -28.0915
17.6535 | -35.1453 | -28.099
17.6596 | -35.1459 | -28.1065
17.6657 | -35.1464 | -28.1141
17.6718 | -35.147 | -28.1217
17.6779 | -35.1475 | -28.1292
17.684 | -35.1481 | -28.1369
17.6901 | -35.1486 | -28.1445
17.6963 | -35.1491 | -28.152
17.7024 | -35.1497 | -28.1595
17.7085 | -35.1502 | -28.1669
17.7146 | -35.1508 | -28.1744
17.7207 | -35.1513 | -28.1818
17.7268 | -35.1519 | -28.1893
17.7329 | -35.1524 | -28.1969
17.739 | -35.153 | -28.2045
17.7451 | -35.1535 | -28.2118
17.7512 | -35.1541 | -28.2194
17.7573 | -35.1546 | -28.2269
17.7634 | -35.1552 | -28.2343
17.7695 | -35.1557 | -28.2418
17.7756 | -35.1563 | -28.2494
17.7817 | -35.1568 | -28.257
17.7878 | -35.1574 | -28.2646
17.7939 | -35.1579 | -28.2722
17.8 | -35.1585 | -28.2795
17.8061 | -35.159 | -28.287
17.8122 | -35.1596 | -28.2945
17.8183 | -35.1601 | -28.302
17.8244 | -35.1607 | -28.3095
17.8305 | -35.1612 | -28.317
17.8366 | -35.1618 | -28.3243
17.8428 | -35.1623 | -28.332
17.8489 | -35.1628 | -28.3395
17.855 | -35.1634 | -28.347
17.8611 | -35.1639 | -28.3544
17.8672 | -35.1645 | -28.3617
17.8733 | -35.1651 | -28.3692
17.8794 | -35.1656 | -28.3766
17.8855 | -35.1662 | -28.3842
17.8916 | -35.1667 | -28.3917
17.8977 | -35.1672 | -28.3992
17.9038 | -35.1678 | -28.4064
17.9099 | -35.1683 | -28.4138
17.916 | -35.1689 | -28.4214
17.9221 | -35.1695 | -28.429
17.9282 | -35.17 | -28.4366
17.9343 | -35.1706 | -28.4441
17.9404 | -35.1711 | -28.4514
17.9465 | -35.1716 | -28.459
17.9526 | -35.1722 | -28.4664
17.9587 | -35.1727 | -28.4739
17.9648 | -35.1733 | -28.4813
17.9709 | -35.1738 | -28.4888
17.977 | -35.1743 | -28.4961
17.9832 | -35.1749 | -28.5036
17.9893 | -35.1754 | -28.5112
17.9954 | -35.176 | -28.5187
18.0015 | -35.1765 | -28.5263
18.0076 | -35.177 | -28.5335
18.0137 | -35.1776 | -28.541
18.0198 | -35.1781 | -28.5484
18.0259 | -35.1786 | -28.5559
18.032 | -35.1792 | -28.5634
18.0381 | -35.1797 | -28.5707
18.0442 | -35.1803 | -28.578
18.0503 | -35.1808 | -28.5855
18.0564 | -35.1814 | -28.5929
18.0625 | -35.1819 | -28.6006
18.0686 | -35.1824 | -28.6082
18.0747 | -35.183 | -28.6157
18.0808 | -35.1836 | -28.623
18.0869 | -35.1841 | -28.6306
18.093 | -35.1847 | -28.6381
18.0991 | -35.1852 | -28.6455
18.1052 | -35.1857 | -28.6531
18.1113 | -35.1862 | -28.6604
18.1174 | -35.1868 | -28.6677
18.1236 | -35.1873 | -28.6751
18.1297 | -35.1879 | -28.6825
18.1358 | -35.1884 | -28.6901
18.1419 | -35.189 | -28.6976
18.148 | -35.1896 | -28.7049
18.1541 | -35.1901 | -28.7121
18.1602 | -35.1907 | -28.7196
18.1663 | -35.1912 | -28.727
18.1724 | -35.1918 | -28.7346
18.1785 | -35.1923 | -28.742
18.1846 | -35.1928 | -28.7492
18.1907 | -35.1934 | -28.7566
18.1968 | -35.1939 | -28.7642
18.2029 | -35.1944 | -28.7716
18.209 | -35.195 | -28.7792
18.2151 | -35.1955 | -28.7868
18.2212 | -35.1961 | -28.7941
18.2273 | -35.1966 | -28.8013
18.2334 | -35.1972 | -28.8089
18.2395 | -35.1977 | -28.8162
18.2456 | -35.1983 | -28.8237
18.2517 | -35.1988 | -28.8312
18.2578 | -35.1994 | -28.8385
18.2639 | -35.1999 | -28.846
18.2701 | -35.2005 | -28.8532
18.2762 | -35.201 | -28.8608
18.2823 | -35.2016 | -28.8679
18.2884 | -35.2022 | -28.8752
18.2945 | -35.2027 | -28.8823
18.3006 | -35.2033 | -28.8898
18.3067 | -35.2038 | -28.8972
18.3128 | -35.2044 | -28.9046
18.3189 | -35.2049 | -28.912
18.325 | -35.2055 | -28.9192
18.3311 | -35.206 | -28.9265
18.3372 | -35.2066 | -28.934
18.3433 | -35.2071 | -28.9415
18.3494 | -35.2077 | -28.9489
18.3555 | -35.2082 | -28.9563
18.3616 | -35.2088 | -28.9635
18.3677 | -35.2093 | -28.9712
18.3738 | -35.2099 | -28.9784
18.3799 | -35.2104 | -28.9857
18.386 | -35.2109 | -28.9931
18.3921 | -35.2114 | -29.0006
18.3982 | -35.212 | -29.0077
18.4043 | -35.2125 | -29.0153
18.4105 | -35.2131 | -29.0226
18.4166 | -35.2136 | -29.03
18.4227 | -35.2141 | -29.0374
18.4288 | -35.2147 | -29.0449
18.4349 | -35.2152 | -29.0522
18.441 | -35.2158 | -29.0595
18.4471 | -35.2163 | -29.0668
18.4532 | -35.2168 | -29.0742
18.4593 | -35.2174 | -29.0816
18.4654 | -35.2179 | -29.0887
18.4715 | -35.2184 | -29.0962
18.4776 | -35.219 | -29.1035
18.4837 | -35.2195 | -29.1111
18.4898 | -35.2201 | -29.1186
18.4959 | -35.2206 | -29.1262
18.502 | -35.2211 | -29.1335
18.5081 | -35.2217 | -29.1411
18.5142 | -35.2222 | -29.1485
18.5203 | -35.2227 | -29.1558
18.5264 | -35.2233 | -29.1631
18.5325 | -35.2238 | -29.1706
18.5386 | -35.2244 | -29.1779
18.5447 | -35.2249 | -29.1853
18.5508 | -35.2254 | -29.1926
18.557 | -35.226 | -29.2
18.5631 | -35.2265 | -29.2073
18.5692 | -35.227 | -29.2146
18.5753 | -35.2275 | -29.2218
18.5814 | -35.2281 | -29.229
18.5875 | -35.2286 | -29.2362
18.5936 | -35.2291 | -29.2435
18.5997 | -35.2296 | -29.251
18.6058 | -35.2302 | -29.2584
18.6119 | -35.2307 | -29.2657
18.618 | -35.2312 | -29.273
18.6241 | -35.2318 | -29.2804
18.6302 | -35.2323 | -29.2879
18.6363 | -35.2328 | -29.2952
18.6424 | -35.2333 | -29.3024
18.6485 | -35.2339 | -29.3098
18.6546 | -35.2344 | -29.3171
18.6607 | -35.2349 | -29.3245
18.6668 | -35.2355 | -29.3319
18.6729 | -35.236 | -29.3395
18.679 | -35.2366 | -29.3466
18.6851 | -35.2371 | -29.3539
18.6912 | -35.2376 | -29.3611
18.6974 | -35.2382 | -29.3684
18.7035 | -35.2387 | -29.3758
18.7096 | -35.2392 | -29.3831
18.7157 | -35.2398 | -29.3903
18.7218 | -35.2403 | -29.3977
18.7279 | -35.2409 | -29.4051
18.734 | -35.2414 | -29.4126
18.7401 | -35.242 | -29.4198
18.7462 | -35.2425 | -29.4269
18.7523 | -35.2431 | -29.4344
18.7584 | -35.2436 | -29.4416
18.7645 | -35.2442 | -29.4489
18.7706 | -35.2447 | -29.4562
18.7767 | -35.2453 | -29.4636
18.7828 | -35.2458 | -29.4707
18.7889 | -35.2463 | -29.4779
18.795 | -35.2469 | -29.4853
18.8011 | -35.2474 | -29.4927
18.8072 | -35.2479 | -29.5003
18.8133 | -35.2484 | -29.5075
18.8194 | -35.2489 | -29.5147
18.8255 | -35.2495 | -29.5219
18.8316 | -35.25 | -29.5294
18.8377 | -35.2505 | -29.5366
18.8439 | -35.2511 | -29.5437
18.85 | -35.2516 | -29.5513
18.8561 | -35.2521 | -29.5584
18.8622 | -35.2527 | -29.5658
18.8683 | -35.2532 | -29.573
18.8744 | -35.2537 | -29.5803
18.8805 | -35.2543 | -29.5877
18.8866 | -35.2548 | -29.5949
18.8927 | -35.2554 | -29.6021
18.8988 | -35.2559 | -29.6092
18.9049 | -35.2564 | -29.6164
18.911 | -35.257 | -29.6238
18.9171 | -35.2575 | -29.631
18.9232 | -35.258 | -29.6383
18.9293 | -35.2586 | -29.6456
18.9354 | -35.2591 | -29.6526
18.9415 | -35.2597 | -29.6599
18.9476 | -35.2602 | -29.6671
18.9537 | -35.2607 | -29.6743
18.9598 | -35.2613 | -29.6813
18.9659 | -35.2618 | -29.6886
18.972 | -35.2623 | -29.6957
18.9781 | -35.2629 | -29.7033
18.9843 | -35.2634 | -29.7106
18.9904 | -35.264 | -29.7179
18.9965 | -35.2645 | -29.7249
19.0026 | -35.2651 | -29.7322
19.0087 | -35.2656 | -29.7396
19.0148 | -35.2662 | -29.7469
19.0209 | -35.2667 | -29.7543
19.027 | -35.2673 | -29.7617
19.0331 | -35.2678 | -29.7687
19.0392 | -35.2683 | -29.776
19.0453 | -35.2689 | -29.7829
19.0514 | -35.2694 | -29.7902
19.0575 | -35.2699 | -29.7974
19.0636 | -35.2705 | -29.8042
19.0697 | -35.2711 | -29.8114
19.0758 | -35.2716 | -29.8188
19.0819 | -35.2722 | -29.8261
19.088 | -35.2727 | -29.8333
19.0941 | -35.2732 | -29.8406
19.1002 | -35.2737 | -29.8477
19.1063 | -35.2742 | -29.8552
19.1124 | -35.2747 | -29.8625
19.1185 | -35.2752 | -29.8697
19.1246 | -35.2757 | -29.877
19.1308 | -35.2763 | -29.8846
19.1369 | -35.2768 | -29.8918
19.143 | -35.2774 | -29.899
19.1491 | -35.2779 | -29.9066
19.1552 | -35.2785 | -29.9139
19.1613 | -35.279 | -29.9212
19.1674 | -35.2795 | -29.9281
19.1735 | -35.2801 | -29.9355
19.1796 | -35.2806 | -29.9431
19.1857 | -35.2812 | -29.9506
19.1918 | -35.2817 | -29.9579
19.1979 | -35.2822 | -29.9649
19.204 | -35.2828 | -29.9721
19.2101 | -35.2833 | -29.9792
19.2162 | -35.2839 | -29.9863
19.2223 | -35.2844 | -29.9932
19.2284 | -35.2849 | -30.0003
19.2345 | -35.2854 | -30.0077
19.2406 | -35.286 | -30.0146
19.2467 | -35.2865 | -30.0216
19.2528 | -35.287 | -30.029
19.2589 | -35.2876 | -30.036
19.265 | -35.2881 | -30.0433
19.2712 | -35.2886 | -30.05
19.2773 | -35.2892 | -30.0571
19.2834 | -35.2897 | -30.0642
19.2895 | -35.2902 | -30.0716
19.2956 | -35.2908 | -30.0787
19.3017 | -35.2913 | -30.0857
19.3078 | -35.2918 | -30.0931
19.3139 | -35.2924 | -30.1002
19.32 | -35.2929 | -30.1075
19.3261 | -35.2934 | -30.1149
19.3322 | -35.2939 | -30.1219
19.3383 | -35.2945 | -30.1293
19.3444 | -35.295 | -30.1363
19.3505 | -35.2955 | -30.1434
19.3566 | -35.2961 | -30.1505
19.3627 | -35.2966 | -30.1576
19.3688 | -35.2972 | -30.1646
19.3749 | -35.2977 | -30.1717
19.381 | -35.2983 | -30.1786
19.3871 | -35.2988 | -30.1859
19.3932 | -35.2993 | -30.1932
19.3993 | -35.2998 | -30.2002
19.4054 | -35.3004 | -30.2075
19.4115 | -35.3009 | -30.2149
19.4177 | -35.3015 | -30.222
19.4238 | -35.302 | -30.2292
19.4299 | -35.3025 | -30.2365
19.436 | -35.3031 | -30.244
19.4421 | -35.3036 | -30.2508
19.4482 | -35.3041 | -30.2581
19.4543 | -35.3046 | -30.2651
19.4604 | -35.3051 | -30.2723
19.4665 | -35.3057 | -30.2793
19.4726 | -35.3062 | -30.2865
19.4787 | -35.3067 | -30.2937
19.4848 | -35.3073 | -30.3009
19.4909 | -35.3078 | -30.3079
19.497 | -35.3083 | -30.3152
19.5031 | -35.3088 | -30.3222
19.5092 | -35.3093 | -30.3296
19.5153 | -35.3098 | -30.3365
19.5214 | -35.3104 | -30.3433
19.5275 | -35.3109 | -30.3504
19.5336 | -35.3113 | -30.3574
19.5397 | -35.3119 | -30.3642
19.5458 | -35.3124 | -30.3715
19.5519 | -35.313 | -30.3786
19.5581 | -35.3135 | -30.3854
19.5642 | -35.314 | -30.3925
19.5703 | -35.3145 | -30.3995
19.5764 | -35.3151 | -30.4066
19.5825 | -35.3156 | -30.4137
19.5886 | -35.3161 | -30.4206
19.5947 | -35.3167 | -30.4275
19.6008 | -35.3172 | -30.4346
19.6069 | -35.3177 | -30.4418
19.613 | -35.3183 | -30.4491
19.6191 | -35.3188 | -30.4561
19.6252 | -35.3193 | -30.463
19.6313 | -35.3198 | -30.4699
19.6374 | -35.3203 | -30.4771
19.6435 | -35.3209 | -30.4843
19.6496 | -35.3214 | -30.4912
19.6557 | -35.3219 | -30.4984
19.6618 | -35.3224 | -30.505
19.6679 | -35.323 | -30.5121
19.674 | -35.3235 | -30.5191
19.6801 | -35.324 | -30.5261
19.6862 | -35.3246 | -30.5332
19.6923 | -35.3251 | -30.5401
19.6984 | -35.3256 | -30.547
19.7046 | -35.3262 | -30.5542
19.7107 | -35.3267 | -30.5614
19.7168 | -35.3272 | -30.5684
19.7229 | -35.3277 | -30.5755
19.729 | -35.3282 | -30.5828
19.7351 | -35.3287 | -30.5898
19.7412 | -35.3292 | -30.5973
19.7473 | -35.3298 | -30.6044
19.7534 | -35.3303 | -30.6117
19.7595 | -35.3308 | -30.619
19.7656 | -35.3313 | -30.6263
19.7717 | -35.3318 | -30.6331
19.7778 | -35.3324 | -30.6403
19.7839 | -35.3329 | -30.6474
19.79 | -35.3334 | -30.6545
19.7961 | -35.3339 | -30.6615
19.8022 | -35.3344 | -30.6684
19.8083 | -35.335 | -30.6755
19.8144 | -35.3355 | -30.6825
19.8205 | -35.3361 | -30.6893
19.8266 | -35.3366 | -30.6961
19.8327 | -35.3371 | -30.7032
19.8388 | -35.3377 | -30.7099
19.845 | -35.3382 | -30.717
19.8511 | -35.3387 | -30.7241
19.8572 | -35.3392 | -30.7309
19.8633 | -35.3397 | -30.7378
19.8694 | -35.3403 | -30.745
19.8755 | -35.3408 | -30.7518
19.8816 | -35.3413 | -30.7587
19.8877 | -35.3418 | -30.7657
19.8938 | -35.3423 | -30.7726
19.8999 | -35.3428 | -30.7797
19.906 | -35.3434 | -30.7866
19.9121 | -35.3439 | -30.7932
19.9182 | -35.3444 | -30.8003
19.9243 | -35.3449 | -30.8074
19.9304 | -35.3455 | -30.8143
19.9365 | -35.346 | -30.8213
19.9426 | -35.3465 | -30.828
19.9487 | -35.3471 | -30.8351
19.9548 | -35.3476 | -30.8421
19.9609 | -35.3481 | -30.8489
19.967 | -35.3487 | -30.8556
19.9731 | -35.3492 | -30.8626
19.9792 | -35.3497 | -30.8695
19.9853 | -35.3503 | -30.8762
19.9915 | -35.3508 | -30.883
19.9976 | -35.3513 | -30.8901
20.0037 | -35.3518 | -30.897
20.0098 | -35.3524 | -30.904
20.0159 | -35.3529 | -30.9108
20.022 | -35.3534 | -30.9177
20.0281 | -35.3539 | -30.925
20.0342 | -35.3545 | -30.9318
20.0403 | -35.355 | -30.9387
20.0464 | -35.3555 | -30.9458
20.0525 | -35.3561 | -30.9527
20.0586 | -35.3567 | -30.9595
20.0647 | -35.3572 | -30.9664
20.0708 | -35.3577 | -30.9734
20.0769 | -35.3582 | -30.9805
20.083 | -35.3587 | -30.9868
20.0891 | -35.3592 | -30.9941
20.0952 | -35.3598 | -31.0012
20.1013 | -35.3603 | -31.0082
20.1074 | -35.3608 | -31.0152
20.1135 | -35.3614 | -31.0224
20.1196 | -35.3619 | -31.0288
20.1257 | -35.3624 | -31.036
20.1319 | -35.3629 | -31.0431
20.138 | -35.3634 | -31.0501
20.1441 | -35.364 | -31.0572
20.1502 | -35.3645 | -31.0645
20.1563 | -35.365 | -31.0712
20.1624 | -35.3655 | -31.0784
20.1685 | -35.366 | -31.0854
20.1746 | -35.3666 | -31.0923
20.1807 | -35.367 | -31.0994
20.1868 | -35.3676 | -31.1064
20.1929 | -35.3681 | -31.1131
20.199 | -35.3686 | -31.1201
20.2051 | -35.3691 | -31.127
20.2112 | -35.3697 | -31.134
20.2173 | -35.3702 | -31.141
20.2234 | -35.3708 | -31.1481
20.2295 | -35.3713 | -31.1547
20.2356 | -35.3718 | -31.1618
20.2417 | -35.3723 | -31.1687
20.2478 | -35.3728 | -31.1759
20.2539 | -35.3734 | -31.183
20.26 | -35.3739 | -31.1899
20.2661 | -35.3745 | -31.1967
20.2723 | -35.375 | -31.2035
20.2784 | -35.3755 | -31.2105
20.2845 | -35.376 | -31.2173
20.2906 | -35.3766 | -31.2243
20.2967 | -35.3771 | -31.2312
20.3028 | -35.3776 | -31.2381
20.3089 | -35.3782 | -31.2452
20.315 | -35.3787 | -31.252
20.3211 | -35.3792 | -31.2591
20.3272 | -35.3797 | -31.266
20.3333 | -35.3802 | -31.273
20.3394 | -35.3808 | -31.2799
20.3455 | -35.3813 | -31.287
20.3516 | -35.3818 | -31.2937
20.3577 | -35.3823 | -31.3006
20.3638 | -35.3828 | -31.3076
20.3699 | -35.3834 | -31.3143
20.376 | -35.3839 | -31.3212
20.3821 | -35.3844 | -31.3283
20.3882 | -35.385 | -31.335
20.3943 | -35.3855 | -31.3418
20.4004 | -35.386 | -31.3481
20.4065 | -35.3865 | -31.355
20.4126 | -35.387 | -31.3619
20.4188 | -35.3876 | -31.3688
20.4249 | -35.3881 | -31.3756
20.431 | -35.3886 | -31.3824
20.4371 | -35.3891 | -31.3893
20.4432 | -35.3896 | -31.3962
20.4493 | -35.3901 | -31.4032
20.4554 | -35.3906 | -31.41
20.4615 | -35.3912 | -31.4167
20.4676 | -35.3918 | -31.4237
20.4737 | -35.3923 | -31.4302
20.4798 | -35.3929 | -31.4372
20.4859 | -35.3934 | -31.4445
20.492 | -35.3939 | -31.4513
20.4981 | -35.3944 | -31.4582
20.5042 | -35.395 | -31.4652
20.5103 | -35.3955 | -31.4719
20.5164 | -35.3961 | -31.4788
20.5225 | -35.3966 | -31.4858
20.5286 | -35.3971 | -31.4927
20.5347 | -35.3976 | -31.4995
20.5408 | -35.3981 | -31.5061
20.5469 | -35.3986 | -31.513
20.553 | -35.3992 | -31.52
20.5592 | -35.3997 | -31.5268
20.5653 | -35.4003 | -31.5335
20.5714 | -35.4008 | -31.5408
20.5775 | -35.4012 | -31.5475
20.5836 | -35.4017 | -31.5543
20.5897 | -35.4023 | -31.5613
20.5958 | -35.4028 | -31.5683
20.6019 | -35.4033 | -31.5751
20.608 | -35.4039 | -31.5821
20.6141 | -35.4044 | -31.5892
20.6202 | -35.4049 | -31.596
20.6263 | -35.4055 | -31.6026
20.6324 | -35.406 | -31.6095
20.6385 | -35.4065 | -31.6164
20.6446 | -35.407 | -31.6232
20.6507 | -35.4076 | -31.6298
20.6568 | -35.4081 | -31.6369
20.6629 | -35.4086 | -31.6435
20.669 | -35.4091 | -31.6503
20.6751 | -35.4096 | -31.657
20.6812 | -35.4101 | -31.6634
20.6873 | -35.4107 | -31.6703
20.6934 | -35.4112 | -31.6772
20.6995 | -35.4117 | -31.6841
20.7057 | -35.4122 | -31.6908
20.7118 | -35.4128 | -31.6976
20.7179 | -35.4133 | -31.7038
20.724 | -35.4138 | -31.7108
20.7301 | -35.4142 | -31.7173
20.7362 | -35.4147 | -31.7241
20.7423 | -35.4153 | -31.7309
20.7484 | -35.4158 | -31.7376
20.7545 | -35.4164 | -31.7444
20.7606 | -35.4169 | -31.7516
20.7667 | -35.4175 | -31.7582
20.7728 | -35.418 | -31.7651
20.7789 | -35.4186 | -31.7721
20.785 | -35.4191 | -31.7789
20.7911 | -35.4197 | -31.7854
20.7972 | -35.4202 | -31.7921
20.8033 | -35.4207 | -31.7991
20.8094 | -35.4212 | -31.806
20.8155 | -35.4217 | -31.813
20.8216 | -35.4222 | -31.8197
20.8277 | -35.4227 | -31.8262
20.8338 | -35.4233 | -31.8328
20.8399 | -35.4238 | -31.8393
20.8461 | -35.4243 | -31.8463
20.8522 | -35.4248 | -31.8534
20.8583 | -35.4254 | -31.8598
20.8644 | -35.4259 | -31.8667
20.8705 | -35.4265 | -31.8736
20.8766 | -35.427 | -31.8803
20.8827 | -35.4276 | -31.887
20.8888 | -35.428 | -31.8939
20.8949 | -35.4286 | -31.9005
20.901 | -35.4291 | -31.9072
20.9071 | -35.4296 | -31.9138
20.9132 | -35.4302 | -31.9205
20.9193 | -35.4307 | -31.9275
20.9254 | -35.4312 | -31.934
20.9315 | -35.4318 | -31.9408
20.9376 | -35.4323 | -31.9475
20.9437 | -35.4329 | -31.9544
20.9498 | -35.4334 | -31.9611
20.9559 | -35.4339 | -31.9676
20.962 | -35.4344 | -31.9742
20.9681 | -35.435 | -31.9809
20.9742 | -35.4355 | -31.9876
20.9803 | -35.436 | -31.9943
20.9864 | -35.4365 | -32.0013
20.9926 | -35.4371 | -32.008
20.9987 | -35.4376 | -32.0147
21.0048 | -35.4382 | -32.0217
21.0109 | -35.4387 | -32.0284
21.017 | -35.4393 | -32.0352
21.0231 | -35.4398 | -32.0419
21.0292 | -35.4403 | -32.0487
21.0353 | -35.4408 | -32.0551
21.0414 | -35.4414 | -32.0619
21.0475 | -35.4419 | -32.0688
21.0536 | -35.4424 | -32.0754
21.0597 | -35.443 | -32.0822
21.0658 | -35.4435 | -32.0888
21.0719 | -35.444 | -32.0955
21.078 | -35.4445 | -32.1023
21.0841 | -35.445 | -32.109
21.0902 | -35.4456 | -32.1156
21.0963 | -35.4461 | -32.1222
21.1024 | -35.4467 | -32.1289
21.1085 | -35.4472 | -32.1351
21.1146 | -35.4478 | -32.1417
21.1207 | -35.4482 | -32.1488
21.1268 | -35.4487 | -32.1556
21.133 | -35.4493 | -32.1625
21.1391 | -35.4498 | -32.1689
21.1452 | -35.4503 | -32.1755
21.1513 | -35.4509 | -32.1822
21.1574 | -35.4514 | -32.189
21.1635 | -35.452 | -32.1955
21.1696 | -35.4525 | -32.2023
21.1757 | -35.453 | -32.2088
21.1818 | -35.4536 | -32.2155
21.1879 | -35.4541 | -32.2222
21.194 | -35.4546 | -32.229
21.2001 | -35.4552 | -32.2355
21.2062 | -35.4557 | -32.2424
21.2123 | -35.4562 | -32.2488
21.2184 | -35.4568 | -32.2554
21.2245 | -35.4573 | -32.2622
21.2306 | -35.4578 | -32.2689
21.2367 | -35.4583 | -32.2758
21.2428 | -35.4589 | -32.2823
21.2489 | -35.4593 | -32.2887
21.255 | -35.4599 | -32.2953
21.2611 | -35.4604 | -32.302
21.2672 | -35.4609 | -32.3089
21.2733 | -35.4614 | -32.3157
21.2795 | -35.462 | -32.3222
21.2856 | -35.4625 | -32.3287
21.2917 | -35.4631 | -32.3356
21.2978 | -35.4636 | -32.3421
21.3039 | -35.4642 | -32.3487
21.31 | -35.4647 | -32.3556
21.3161 | -35.4652 | -32.3621
21.3222 | -35.4657 | -32.3688
21.3283 | -35.4663 | -32.3755
21.3344 | -35.4668 | -32.3823
21.3405 | -35.4673 | -32.3891
21.3466 | -35.4679 | -32.3958
21.3527 | -35.4684 | -32.4021
21.3588 | -35.469 | -32.4088
21.3649 | -35.4695 | -32.4155
21.371 | -35.47 | -32.4222
21.3771 | -35.4706 | -32.4288
21.3832 | -35.4711 | -32.4355
21.3893 | -35.4716 | -32.4421
21.3954 | -35.4722 | -32.4489
21.4015 | -35.4727 | -32.4555
21.4076 | -35.4732 | -32.4624
21.4137 | -35.4738 | -32.4691
21.4199 | -35.4743 | -32.4757
21.426 | -35.4748 | -32.4819
21.4321 | -35.4753 | -32.4886
21.4382 | -35.4758 | -32.4953
21.4443 | -35.4763 | -32.5019
21.4504 | -35.4769 | -32.5085
21.4565 | -35.4774 | -32.5146
21.4626 | -35.4779 | -32.5213
21.4687 | -35.4784 | -32.5281
21.4748 | -35.4789 | -32.5346
21.4809 | -35.4795 | -32.5411
21.487 | -35.48 | -32.5476
21.4931 | -35.4805 | -32.5538
21.4992 | -35.4811 | -32.5604
21.5053 | -35.4816 | -32.5669
21.5114 | -35.4821 | -32.5738
21.5175 | -35.4826 | -32.5803
21.5236 | -35.4831 | -32.587
21.5297 | -35.4836 | -32.5935
21.5358 | -35.4841 | -32.6001
21.5419 | -35.4847 | -32.607
21.548 | -35.4852 | -32.6138
21.5541 | -35.4857 | -32.6203
21.5602 | -35.4863 | -32.627
21.5664 | -35.4868 | -32.6332
21.5725 | -35.4873 | -32.6398
21.5786 | -35.4879 | -32.6463
21.5847 | -35.4884 | -32.6532
21.5908 | -35.4889 | -32.6598
21.5969 | -35.4895 | -32.666
21.603 | -35.49 | -32.6727
21.6091 | -35.4905 | -32.6795
21.6152 | -35.491 | -32.686
21.6213 | -35.4916 | -32.6927
21.6274 | -35.4921 | -32.6994
21.6335 | -35.4926 | -32.7057
21.6396 | -35.4931 | -32.7124
21.6457 | -35.4937 | -32.7191
21.6518 | -35.4942 | -32.7256
21.6579 | -35.4947 | -32.7322
21.664 | -35.4953 | -32.7387
21.6701 | -35.4958 | -32.7449
21.6762 | -35.4963 | -32.7513
21.6823 | -35.4969 | -32.7579
21.6884 | -35.4974 | -32.7644
21.6945 | -35.4979 | -32.771
21.7006 | -35.4984 | -32.7776
21.7068 | -35.4989 | -32.7839
21.7129 | -35.4994 | -32.7902
21.719 | -35.4999 | -32.7968
21.7251 | -35.5005 | -32.8033
21.7312 | -35.501 | -32.81
21.7373 | -35.5015 | -32.8162
21.7434 | -35.5021 | -32.8228
21.7495 | -35.5026 | -32.8294
21.7556 | -35.5031 | -32.8359
21.7617 | -35.5037 | -32.8425
21.7678 | -35.5042 | -32.8488
21.7739 | -35.5048 | -32.855
21.78 | -35.5053 | -32.8618
21.7861 | -35.5059 | -32.8685
21.7922 | -35.5064 | -32.8751
21.7983 | -35.5069 | -32.8817
21.8044 | -35.5074 | -32.8884
21.8105 | -35.508 | -32.8946
21.8166 | -35.5085 | -32.9013
21.8227 | -35.509 | -32.9078
21.8288 | -35.5095 | -32.9144
21.8349 | -35.5101 | -32.9211
21.841 | -35.5105 | -32.9275
21.8471 | -35.5111 | -32.9337
21.8533 | -35.5116 | -32.9401
21.8594 | -35.5121 | -32.9466
21.8655 | -35.5127 | -32.9532
21.8716 | -35.5132 | -32.9597
21.8777 | -35.5138 | -32.9662
21.8838 | -35.5143 | -32.9727
21.8899 | -35.5148 | -32.9792
21.896 | -35.5154 | -32.9858
21.9021 | -35.5159 | -32.9925
21.9082 | -35.5164 | -32.9992
21.9143 | -35.517 | -33.0056
21.9204 | -35.5175 | -33.0123
21.9265 | -35.5181 | -33.0188
21.9326 | -35.5186 | -33.0252
21.9387 | -35.5191 | -33.0318
21.9448 | -35.5196 | -33.0383
21.9509 | -35.5202 | -33.0445
21.957 | -35.5207 | -33.0512
21.9631 | -35.5212 | -33.0578
21.9692 | -35.5218 | -33.0644
21.9753 | -35.5223 | -33.071
21.9814 | -35.5229 | -33.0776
21.9875 | -35.5234 | -33.0837
21.9937 | -35.5239 | -33.0903
21.9998 | -35.5244 | -33.0969
22.0059 | -35.525 | -33.1034
22.012 | -35.5255 | -33.1098
22.0181 | -35.526 | -33.116
22.0242 | -35.5266 | -33.1225
22.0303 | -35.5271 | -33.1291
22.0364 | -35.5276 | -33.1356
22.0425 | -35.5281 | -33.1422
22.0486 | -35.5287 | -33.1485
22.0547 | -35.5292 | -33.1547
22.0608 | -35.5297 | -33.1608
22.0669 | -35.5303 | -33.1671
22.073 | -35.5308 | -33.1736
22.0791 | -35.5313 | -33.1801
22.0852 | -35.5318 | -33.1868
22.0913 | -35.5323 | -33.1932
22.0974 | -35.5328 | -33.1996
22.1035 | -35.5333 | -33.2062
22.1096 | -35.5339 | -33.2126
22.1157 | -35.5344 | -33.2193
22.1218 | -35.5349 | -33.2256
22.1279 | -35.5354 | -33.2318
22.134 | -35.536 | -33.2383
22.1402 | -35.5365 | -33.2446
22.1463 | -35.537 | -33.251
22.1524 | -35.5375 | -33.2573
22.1585 | -35.5381 | -33.2637
22.1646 | -35.5386 | -33.2698
22.1707 | -35.5391 | -33.2763
22.1768 | -35.5396 | -33.2828
22.1829 | -35.5402 | -33.2893
22.189 | -35.5407 | -33.2958
22.1951 | -35.5412 | -33.3019
22.2012 | -35.5417 | -33.3083
22.2073 | -35.5423 | -33.3146
22.2134 | -35.5428 | -33.321
22.2195 | -35.5433 | -33.3273
22.2256 | -35.5438 | -33.3338
22.2317 | -35.5444 | -33.3397
22.2378 | -35.5449 | -33.3462
22.2439 | -35.5454 | -33.3526
22.25 | -35.5459 | -33.359
22.2561 | -35.5464 | -33.3654
22.2622 | -35.547 | -33.3718
22.2683 | -35.5475 | -33.3781
22.2744 | -35.548 | -33.3845
22.2806 | -35.5486 | -33.3911
22.2867 | -35.5491 | -33.3973
22.2928 | -35.5496 | -33.4036
22.2989 | -35.5502 | -33.41
22.305 | -35.5507 | -33.4161
22.3111 | -35.5512 | -33.4226
22.3172 | -35.5518 | -33.4291
22.3233 | -35.5523 | -33.4355
22.3294 | -35.5529 | -33.442
22.3355 | -35.5534 | -33.4481
22.3416 | -35.5539 | -33.4545
22.3477 | -35.5545 | -33.461
22.3538 | -35.555 | -33.4672
22.3599 | -35.5555 | -33.4737
22.366 | -35.556 | -33.48
22.3721 | -35.5565 | -33.4863
22.3782 | -35.557 | -33.4927
22.3843 | -35.5575 | -33.499
22.3904 | -35.5581 | -33.5054
22.3965 | -35.5586 | -33.512
22.4026 | -35.5591 | -33.5184
22.4087 | -35.5596 | -33.5243
22.4148 | -35.5601 | -33.5307
22.4209 | -35.5607 | -33.537
22.4271 | -35.5612 | -33.5434
22.4332 | -35.5617 | -33.5497
22.4393 | -35.5623 | -33.5562
22.4454 | -35.5628 | -33.5622
22.4515 | -35.5633 | -33.5686
22.4576 | -35.5638 | -33.5751
22.4637 | -35.5643 | -33.5816
22.4698 | -35.5649 | -33.5879
22.4759 | -35.5654 | -33.5941
22.482 | -35.566 | -33.6005
22.4881 | -35.5665 | -33.6069
22.4942 | -35.567 | -33.6133
22.5003 | -35.5675 | -33.6198
22.5064 | -35.568 | -33.6261
22.5125 | -35.5685 | -33.632
22.5186 | -35.5691 | -33.6384
22.5247 | -35.5696 | -33.6447
22.5308 | -35.5701 | -33.651
22.5369 | -35.5706 | -33.6573
22.543 | -35.5711 | -33.6637
22.5491 | -35.5716 | -33.6698
22.5552 | -35.5722 | -33.6762
22.5613 | -35.5727 | -33.6827
22.5675 | -35.5733 | -33.689
22.5736 | -35.5738 | -33.6954
22.5797 | -35.5743 | -33.7017
22.5858 | -35.5749 | -33.7077
22.5919 | -35.5754 | -33.714
22.598 | -35.5759 | -33.7205
22.6041 | -35.5764 | -33.7269
22.6102 | -35.577 | -33.7332
22.6163 | -35.5775 | -33.7393
22.6224 | -35.578 | -33.7456
22.6285 | -35.5785 | -33.7519
22.6346 | -35.579 | -33.7583
22.6407 | -35.5795 | -33.7646
22.6468 | -35.5801 | -33.7709
22.6529 | -35.5806 | -33.7768
22.659 | -35.5811 | -33.7829
22.6651 | -35.5817 | -33.7893
22.6712 | -35.5822 | -33.7956
22.6773 | -35.5827 | -33.802
22.6834 | -35.5832 | -33.8083
22.6895 | -35.5838 | -33.8144
22.6956 | -35.5843 | -33.8206
22.7017 | -35.5849 | -33.8271
22.7079 | -35.5854 | -33.8332
22.714 | -35.5859 | -33.8395
22.7201 | -35.5865 | -33.8458
22.7262 | -35.587 | -33.852
22.7323 | -35.5875 | -33.8583
22.7384 | -35.5881 | -33.8646
22.7445 | -35.5886 | -33.871
22.7506 | -35.5891 | -33.8772
22.7567 | -35.5896 | -33.8836
22.7628 | -35.5901 | -33.8895
22.7689 | -35.5907 | -33.8957
22.775 | -35.5912 | -33.902
22.7811 | -35.5917 | -33.9081
22.7872 | -35.5923 | -33.9145
22.7933 | -35.5928 | -33.9204
22.7994 | -35.5933 | -33.9268
22.8055 | -35.5938 | -33.933
22.8116 | -35.5943 | -33.9393
22.8177 | -35.5949 | -33.9457
22.8238 | -35.5954 | -33.9519
22.8299 | -35.596 | -33.9581
22.836 | -35.5965 | -33.9643
22.8421 | -35.597 | -33.9707
22.8482 | -35.5975 | -33.977
22.8544 | -35.598 | -33.9834
22.8605 | -35.5986 | -33.9896
22.8666 | -35.5991 | -33.9955
22.8727 | -35.5996 | -34.0017
22.8788 | -35.6001 | -34.0078
22.8849 | -35.6006 | -34.0139
22.891 | -35.6011 | -34.0202
22.8971 | -35.6017 | -34.0264
22.9032 | -35.6022 | -34.0324
22.9093 | -35.6027 | -34.0386
22.9154 | -35.6032 | -34.0449
22.9215 | -35.6037 | -34.0511
22.9276 | -35.6043 | -34.0574
22.9337 | -35.6048 | -34.0633
22.9398 | -35.6053 | -34.0696
22.9459 | -35.6058 | -34.076
22.952 | -35.6063 | -34.0821
22.9581 | -35.6068 | -34.0883
22.9642 | -35.6074 | -34.0947
22.9703 | -35.6079 | -34.1005
22.9764 | -35.6084 | -34.1067
22.9825 | -35.6089 | -34.113
22.9886 | -35.6094 | -34.1193
22.9948 | -35.6099 | -34.1255
23.0009 | -35.6105 | -34.1319
23.007 | -35.611 | -34.138
23.0131 | -35.6115 | -34.1441
23.0192 | -35.612 | -34.1504
23.0253 | -35.6125 | -34.1567
23.0314 | -35.613 | -34.1629
23.0375 | -35.6135 | -34.1691
23.0436 | -35.614 | -34.1752
23.0497 | -35.6146 | -34.1815
23.0558 | -35.6151 | -34.1877
23.0619 | -35.6156 | -34.1939
23.068 | -35.6161 | -34.2002
23.0741 | -35.6166 | -34.2062
23.0802 | -35.6171 | -34.2122
23.0863 | -35.6176 | -34.2184
23.0924 | -35.6181 | -34.2245
23.0985 | -35.6187 | -34.2306
23.1046 | -35.6192 | -34.2369
23.1107 | -35.6197 | -34.243
23.1168 | -35.6202 | -34.2493
23.1229 | -35.6208 | -34.2554
23.129 | -35.6213 | -34.2615
23.1351 | -35.6218 | -34.2677
23.1413 | -35.6223 | -34.2738
23.1474 | -35.6228 | -34.2798
23.1535 | -35.6233 | -34.286
23.1596 | -35.6239 | -34.2922
23.1657 | -35.6244 | -34.2984
23.1718 | -35.6249 | -34.3046
23.1779 | -35.6255 | -34.311
23.184 | -35.626 | -34.3169
23.1901 | -35.6265 | -34.3231
23.1962 | -35.627 | -34.3294
23.2023 | -35.6275 | -34.3355
23.2084 | -35.628 | -34.3417
23.2145 | -35.6285 | -34.3473
23.2206 | -35.629 | -34.3535
23.2267 | -35.6296 | -34.3596
23.2328 | -35.6301 | -34.3659
23.2389 | -35.6306 | -34.372
23.245 | -35.6311 | -34.3781
23.2511 | -35.6316 | -34.3841
23.2572 | -35.6321 | -34.3904
23.2633 | -35.6326 | -34.3966
23.2694 | -35.6331 | -34.4029
23.2755 | -35.6337 | -34.4093
23.2817 | -35.6342 | -34.4154
23.2878 | -35.6347 | -34.4213
23.2939 | -35.6352 | -34.4275
23.3 | -35.6357 | -34.4337
23.3061 | -35.6362 | -34.4397
23.3122 | -35.6367 | -34.4458
23.3183 | -35.6372 | -34.4518
23.3244 | -35.6377 | -34.4576
23.3305 | -35.6382 | -34.4637
23.3366 | -35.6387 | -34.4699
23.3427 | -35.6393 | -34.4759
23.3488 | -35.6398 | -34.4821
23.3549 | -35.6403 | -34.4882
23.361 | -35.6408 | -34.494
23.3671 | -35.6413 | -34.5002
23.3732 | -35.6418 | -34.5064
23.3793 | -35.6424 | -34.5125
23.3854 | -35.6429 | -34.5189
23.3915 | -35.6434 | -34.5249
23.3976 | -35.6439 | -34.531
23.4037 | -35.6444 | -34.5372
23.4098 | -35.645 | -34.5434
23.4159 | -35.6455 | -34.5496
23.422 | -35.646 | -34.5556
23.4282 | -35.6465 | -34.5614
23.4343 | -35.647 | -34.5676
23.4404 | -35.6475 | -34.5736
23.4465 | -35.648 | -34.5796
23.4526 | -35.6485 | -34.5859
23.4587 | -35.649 | -34.5918
23.4648 | -35.6495 | -34.5975
23.4709 | -35.65 | -34.6036
23.477 | -35.6505 | -34.6098
23.4831 | -35.651 | -34.616
23.4892 | -35.6515 | -34.6221
23.4953 | -35.652 | -34.6282
23.5014 | -35.6525 | -34.6341
23.5075 | -35.653 | -34.6402
23.5136 | -35.6536 | -34.6463
23.5197 | -35.6541 | -34.6523
23.5258 | -35.6546 | -34.6587
23.5319 | -35.6551 | -34.6644
23.538 | -35.6555 | -34.6704
23.5441 | -35.6561 | -34.6764
23.5502 | -35.6566 | -34.6825
23.5563 | -35.6571 | -34.6887
23.5624 | -35.6576 | -34.6948
23.5686 | -35.6581 | -34.7006
23.5747 | -35.6586 | -34.7065
23.5808 | -35.6592 | -34.7124
23.5869 | -35.6597 | -34.7185
23.593 | -35.6602 | -34.7246
23.5991 | -35.6607 | -34.7308
23.6052 | -35.6612 | -34.7364
23.6113 | -35.6617 | -34.7427
23.6174 | -35.6622 | -34.7488
23.6235 | -35.6627 | -34.7548
23.6296 | -35.6632 | -34.7609
23.6357 | -35.6637 | -34.7671
23.6418 | -35.6642 | -34.7728
23.6479 | -35.6647 | -34.7788
23.654 | -35.6652 | -34.7849
23.6601 | -35.6657 | -34.7909
23.6662 | -35.6662 | -34.7969
23.6723 | -35.6667 | -34.8026
23.6784 | -35.6672 | -34.8086
23.6845 | -35.6677 | -34.8145
23.6906 | -35.6683 | -34.8204
23.6967 | -35.6688 | -34.8266
23.7028 | -35.6693 | -34.8325
23.7089 | -35.6698 | -34.8384
23.7151 | -35.6703 | -34.8444
23.7212 | -35.6708 | -34.8505
23.7273 | -35.6712 | -34.8566
23.7334 | -35.6718 | -34.8627
23.7395 | -35.6723 | -34.8689
23.7456 | -35.6728 | -34.8747
23.7517 | -35.6733 | -34.8808
23.7578 | -35.6738 | -34.8866
23.7639 | -35.6743 | -34.8927
23.77 | -35.6748 | -34.8988
23.7761 | -35.6753 | -34.9049
23.7822 | -35.6758 | -34.9105
23.7883 | -35.6763 | -34.9163
23.7944 | -35.6768 | -34.9223
23.8005 | -35.6773 | -34.9285
23.8066 | -35.6778 | -34.9345
23.8127 | -35.6783 | -34.9403
23.8188 | -35.6788 | -34.9464
23.8249 | -35.6793 | -34.9524
23.831 | -35.6798 | -34.9586
23.8371 | -35.6803 | -34.9647
23.8432 | -35.6808 | -34.9708
23.8493 | -35.6813 | -34.9766
23.8555 | -35.6818 | -34.9825
23.8616 | -35.6823 | -34.9884
23.8677 | -35.6828 | -34.9945
23.8738 | -35.6833 | -35.0003
23.8799 | -35.6838 | -35.0063
23.886 | -35.6842 | -35.0122
23.8921 | -35.6847 | -35.0183
23.8982 | -35.6852 | -35.0242
23.9043 | -35.6857 | -35.0302
23.9104 | -35.6861 | -35.0361
23.9165 | -35.6866 | -35.0422
23.9226 | -35.6871 | -35.048
23.9287 | -35.6876 | -35.054
23.9348 | -35.6881 | -35.06
23.9409 | -35.6886 | -35.0661
23.947 | -35.6891 | -35.0723
23.9531 | -35.6896 | -35.0778
23.9592 | -35.6901 | -35.0837
23.9653 | -35.6906 | -35.0896
23.9714 | -35.6911 | -35.0956
23.9775 | -35.6916 | -35.1016
23.9836 | -35.6921 | -35.1075
23.9897 | -35.6925 | -35.1133
23.9958 | -35.693 | -35.1193
24.002 | -35.6935 | -35.1255
24.0081 | -35.694 | -35.1314
24.0142 | -35.6945 | -35.1374
24.0203 | -35.695 | -35.1435
24.0264 | -35.6955 | -35.1493
24.0325 | -35.696 | -35.1552
24.0386 | -35.6965 | -35.1613
24.0447 | -35.6969 | -35.1672
24.0508 | -35.6975 | -35.1731
24.0569 | -35.698 | -35.1791
24.063 | -35.6985 | -35.1847
24.0691 | -35.6989 | -35.1906
24.0752 | -35.6994 | -35.1968
24.0813 | -35.6999 | -35.2028
24.0874 | -35.7004 | -35.2088
24.0935 | -35.7009 | -35.2148
24.0996 | -35.7014 | -35.2205
24.1057 | -35.7019 | -35.2264
24.1118 | -35.7024 | -35.2324
24.1179 | -35.7029 | -35.2384
24.124 | -35.7033 | -35.2442
24.1301 | -35.7038 | -35.2498
24.1362 | -35.7043 | -35.2558
24.1424 | -35.7048 | -35.2617
24.1485 | -35.7054 | -35.2676
24.1546 | -35.7059 | -35.2736
24.1607 | -35.7063 | -35.2796
24.1668 | -35.7068 | -35.2853
24.1729 | -35.7073 | -35.2913
24.179 | -35.7078 | -35.2971
24.1851 | -35.7083 | -35.3031
24.1912 | -35.7088 | -35.3088
24.1973 | -35.7093 | -35.3149
24.2034 | -35.7098 | -35.3205
24.2095 | -35.7103 | -35.3265
24.2156 | -35.7108 | -35.3322
24.2217 | -35.7113 | -35.3381
24.2278 | -35.7117 | -35.3441
24.2339 | -35.7122 | -35.3499
24.24 | -35.7127 | -35.3555
24.2461 | -35.7132 | -35.3612
24.2522 | -35.7136 | -35.3671
24.2583 | -35.7141 | -35.3729
24.2644 | -35.7146 | -35.3789
24.2705 | -35.7151 | -35.3847
24.2766 | -35.7156 | -35.3906
24.2827 | -35.716 | -35.3966
24.2889 | -35.7165 | -35.4027
24.295 | -35.717 | -35.4085
24.3011 | -35.7175 | -35.4145
24.3072 | -35.7179 | -35.4203
24.3133 | -35.7184 | -35.4263
24.3194 | -35.7189 | -35.432
24.3255 | -35.7194 | -35.4382
24.3316 | -35.7199 | -35.4442
24.3377 | -35.7204 | -35.45
24.3438 | -35.7208 | -35.4559
24.3499 | -35.7213 | -35.4616
24.356 | -35.7218 | -35.4675
24.3621 | -35.7222 | -35.4735
24.3682 | -35.7228 | -35.4794
24.3743 | -35.7232 | -35.4855
24.3804 | -35.7237 | -35.4912
24.3865 | -35.7242 | -35.4972
24.3926 | -35.7247 | -35.5032
24.3987 | -35.7251 | -35.509
24.4048 | -35.7256 | -35.5148
24.4109 | -35.7261 | -35.5207
24.417 | -35.7266 | -35.5265
24.4231 | -35.7271 | -35.5323
24.4293 | -35.7276 | -35.5381
24.4354 | -35.7281 | -35.544
24.4415 | -35.7286 | -35.5499
24.4476 | -35.729 | -35.5558
24.4537 | -35.7295 | -35.5617
24.4598 | -35.73 | -35.5675
24.4659 | -35.7305 | -35.5734
24.472 | -35.7309 | -35.5792
24.4781 | -35.7314 | -35.5851
24.4842 | -35.7318 | -35.5907
24.4903 | -35.7323 | -35.5965
24.4964 | -35.7328 | -35.6023
24.5025 | -35.7332 | -35.6082
24.5086 | -35.7337 | -35.6141
24.5147 | -35.7342 | -35.6201
24.5208 | -35.7347 | -35.6258
24.5269 | -35.7351 | -35.6317
24.533 | -35.7356 | -35.6376
24.5391 | -35.7362 | -35.6432
24.5452 | -35.7366 | -35.6491
24.5513 | -35.7371 | -35.6547
24.5574 | -35.7376 | -35.6606
24.5635 | -35.738 | -35.6665
24.5696 | -35.7385 | -35.6723
24.5758 | -35.739 | -35.6782
24.5819 | -35.7395 | -35.6839
24.588 | -35.7399 | -35.6897
24.5941 | -35.7404 | -35.6955
24.6002 | -35.7409 | -35.7013
24.6063 | -35.7414 | -35.7072
24.6124 | -35.7418 | -35.7131
24.6185 | -35.7423 | -35.7188
24.6246 | -35.7428 | -35.7245
24.6307 | -35.7433 | -35.7304
24.6368 | -35.7438 | -35.7363
24.6429 | -35.7442 | -35.7421
24.649 | -35.7447 | -35.7481
24.6551 | -35.7452 | -35.7538
24.6612 | -35.7456 | -35.7592
24.6673 | -35.7461 | -35.765
24.6734 | -35.7466 | -35.7707
24.6795 | -35.7471 | -35.7764
24.6856 | -35.7475 | -35.7824
24.6917 | -35.748 | -35.7882
24.6978 | -35.7484 | -35.7938
24.7039 | -35.7489 | -35.7998
24.71 | -35.7494 | -35.8056
24.7162 | -35.7498 | -35.8115
24.7223 | -35.7503 | -35.8173
24.7284 | -35.7507 | -35.8228
24.7345 | -35.7512 | -35.8287
24.7406 | -35.7517 | -35.8346
24.7467 | -35.7521 | -35.8406
24.7528 | -35.7526 | -35.8464
24.7589 | -35.7531 | -35.8522
24.765 | -35.7535 | -35.8579
24.7711 | -35.754 | -35.8637
24.7772 | -35.7545 | -35.8694
24.7833 | -35.7549 | -35.8752
24.7894 | -35.7554 | -35.8811
24.7955 | -35.7559 | -35.8871
24.8016 | -35.7564 | -35.8928
24.8077 | -35.7569 | -35.8986
24.8138 | -35.7573 | -35.9044
24.8199 | -35.7578 | -35.9103
24.826 | -35.7583 | -35.9161
24.8321 | -35.7588 | -35.9219
24.8382 | -35.7593 | -35.9272
24.8443 | -35.7597 | -35.9329
24.8504 | -35.7601 | -35.9386
24.8565 | -35.7606 | -35.9444
24.8627 | -35.7611 | -35.9501
24.8688 | -35.7615 | -35.9558
24.8749 | -35.762 | -35.9616
24.881 | -35.7625 | -35.9673
24.8871 | -35.7629 | -35.9732
24.8932 | -35.7634 | -35.9791
24.8993 | -35.7639 | -35.985
24.9054 | -35.7643 | -35.9906
24.9115 | -35.7648 | -35.9963
24.9176 | -35.7653 | -36.0021
24.9237 | -35.7658 | -36.0079
24.9298 | -35.7662 | -36.0138
24.9359 | -35.7667 | -36.0197
24.942 | -35.7671 | -36.0254
24.9481 | -35.7675 | -36.0313
24.9542 | -35.768 | -36.037
24.9603 | -35.7685 | -36.0428
24.9664 | -35.7689 | -36.0484
24.9725 | -35.7694 | -36.0541
24.9786 | -35.7698 | -36.0596
24.9847 | -35.7703 | -36.0653
24.9908 | -35.7707 | -36.0711
24.9969 | -35.7712 | -36.0769
25.0031 | -35.7717 | -36.0828
25.0092 | -35.7721 | -36.0884
25.0153 | -35.7725 | -36.0941
25.0214 | -35.773 | -36.0998
25.0275 | -35.7735 | -36.1055
25.0336 | -35.7739 | -36.1112
25.0397 | -35.7744 | -36.1169
25.0458 | -35.7749 | -36.1226
25.0519 | -35.7753 | -36.1283
25.058 | -35.7757 | -36.1339
25.0641 | -35.7762 | -36.1395
25.0702 | -35.7766 | -36.1451
25.0763 | -35.7771 | -36.151
25.0824 | -35.7775 | -36.1566
25.0885 | -35.7779 | -36.1624
25.0946 | -35.7784 | -36.168
25.1007 | -35.7788 | -36.1738
25.1068 | -35.7793 | -36.1796
25.1129 | -35.7798 | -36.1853
25.119 | -35.7802 | -36.1908
25.1251 | -35.7807 | -36.1965
25.1312 | -35.7812 | -36.2023
25.1373 | -35.7816 | -36.208
25.1435 | -35.7821 | -36.2139
25.1496 | -35.7825 | -36.2195
25.1557 | -35.783 | -36.2254
25.1618 | -35.7834 | -36.231
25.1679 | -35.7839 | -36.2368
25.174 | -35.7843 | -36.2424
25.1801 | -35.7848 | -36.2484
25.1862 | -35.7853 | -36.2541
25.1923 | -35.7857 | -36.2599
25.1984 | -35.7862 | -36.2655
25.2045 | -35.7866 | -36.2713
25.2106 | -35.7871 | -36.2772
25.2167 | -35.7875 | -36.283
25.2228 | -35.7879 | -36.2885
25.2289 | -35.7884 | -36.294
25.235 | -35.7888 | -36.2997
25.2411 | -35.7893 | -36.3054
25.2472 | -35.7897 | -36.3112
25.2533 | -35.7902 | -36.3168
25.2594 | -35.7906 | -36.3221
25.2655 | -35.7911 | -36.3277
25.2716 | -35.7915 | -36.3333
25.2777 | -35.792 | -36.339
25.2838 | -35.7924 | -36.3447
25.29 | -35.7929 | -36.3504
25.2961 | -35.7933 | -36.356
25.3022 | -35.7938 | -36.3618
25.3083 | -35.7943 | -36.3674
25.3144 | -35.7947 | -36.3731
25.3205 | -35.7952 | -36.3789
25.3266 | -35.7956 | -36.3843
25.3327 | -35.7961 | -36.3901
25.3388 | -35.7965 | -36.3958
25.3449 | -35.797 | -36.4015
25.351 | -35.7974 | -36.4071
25.3571 | -35.7979 | -36.4129
25.3632 | -35.7983 | -36.4184
25.3693 | -35.7987 | -36.4242
25.3754 | -35.7992 | -36.4299
25.3815 | -35.7996 | -36.4357
25.3876 | -35.8001 | -36.4415
25.3937 | -35.8005 | -36.4471
25.3998 | -35.8009 | -36.4526
25.4059 | -35.8014 | -36.4584
25.412 | -35.8018 | -36.4639
25.4181 | -35.8023 | -36.4696
25.4242 | -35.8027 | -36.4753
25.4304 | -35.8032 | -36.4811
25.4365 | -35.8036 | -36.4868
25.4426 | -35.804 | -36.4925
25.4487 | -35.8044 | -36.4982
25.4548 | -35.8049 | -36.5038
25.4609 | -35.8053 | -36.5096
25.467 | -35.8057 | -36.5152
25.4731 | -35.8062 | -36.5208
25.4792 | -35.8067 | -36.5264
25.4853 | -35.8071 | -36.5321
25.4914 | -35.8076 | -36.5377
25.4975 | -35.808 | -36.5434
25.5036 | -35.8084 | -36.5489
25.5097 | -35.8089 | -36.5545
25.5158 | -35.8093 | -36.5603
25.5219 | -35.8098 | -36.566
25.528 | -35.8102 | -36.5716
25.5341 | -35.8107 | -36.5774
25.5402 | -35.8111 | -36.5828
25.5463 | -35.8115 | -36.5884
25.5524 | -35.812 | -36.594
25.5585 | -35.8124 | -36.5996
25.5646 | -35.8128 | -36.6053
25.5707 | -35.8133 | -36.6111
25.5769 | -35.8137 | -36.6164
25.583 | -35.8141 | -36.6222
25.5891 | -35.8146 | -36.628
25.5952 | -35.815 | -36.6338
25.6013 | -35.8155 | -36.6394
25.6074 | -35.8159 | -36.6448
25.6135 | -35.8163 | -36.6506
25.6196 | -35.8168 | -36.6563
25.6257 | -35.8172 | -36.6621
25.6318 | -35.8177 | -36.6677
25.6379 | -35.8181 | -36.6735
25.644 | -35.8185 | -36.679
25.6501 | -35.819 | -36.6846
25.6562 | -35.8194 | -36.6903
25.6623 | -35.8198 | -36.6957
25.6684 | -35.8203 | -36.7014
25.6745 | -35.8208 | -36.7071
25.6806 | -35.8212 | -36.7126
25.6867 | -35.8216 | -36.7182
25.6928 | -35.822 | -36.7238
25.6989 | -35.8225 | -36.7294
25.705 | -35.8229 | -36.735
25.7111 | -35.8234 | -36.7407
25.7173 | -35.8238 | -36.7462
25.7234 | -35.8243 | -36.7519
25.7295 | -35.8247 | -36.7575
25.7356 | -35.8251 | -36.7634
25.7417 | -35.8255 | -36.7689
25.7478 | -35.826 | -36.7744
25.7539 | -35.8264 | -36.78
25.76 | -35.8268 | -36.7855
25.7661 | -35.8273 | -36.7911
25.7722 | -35.8277 | -36.7968
25.7783 | -35.8281 | -36.8025
25.7844 | -35.8285 | -36.8079
25.7905 | -35.829 | -36.8136
25.7966 | -35.8294 | -36.8192
25.8027 | -35.8298 | -36.825
25.8088 | -35.8302 | -36.8309
25.8149 | -35.8307 | -36.8366
25.821 | -35.8311 | -36.8419
25.8271 | -35.8315 | -36.8475
25.8332 | -35.8319 | -36.8533
25.8393 | -35.8324 | -36.8589
25.8454 | -35.8328 | -36.8645
25.8515 | -35.8332 | -36.8702
25.8576 | -35.8336 | -36.8756
25.8638 | -35.8341 | -36.8812
25.8699 | -35.8345 | -36.8868
25.876 | -35.8349 | -36.8922
25.8821 | -35.8353 | -36.8979
25.8882 | -35.8358 | -36.9032
25.8943 | -35.8362 | -36.909
25.9004 | -35.8366 | -36.9147
25.9065 | -35.837 | -36.9204
25.9126 | -35.8375 | -36.9261
25.9187 | -35.8379 | -36.9318
25.9248 | -35.8383 | -36.9372
25.9309 | -35.8387 | -36.9429
25.937 | -35.8392 | -36.9486
25.9431 | -35.8396 | -36.9545
25.9492 | -35.84 | -36.96
25.9553 | -35.8405 | -36.9657
25.9614 | -35.8409 | -36.9712
25.9675 | -35.8413 | -36.9768
25.9736 | -35.8417 | -36.9827
25.9797 | -35.8422 | -36.9883
25.9858 | -35.8426 | -36.994
25.9919 | -35.843 | -36.9997
25.998 | -35.8434 | -37.005
26.0042 | -35.8438 | -37.0108
26.0103 | -35.8442 | -37.0163
26.0164 | -35.8446 | -37.0219
26.0225 | -35.8451 | -37.0275
26.0286 | -35.8455 | -37.0331
26.0347 | -35.8459 | -37.0384
26.0408 | -35.8463 | -37.0441
26.0469 | -35.8467 | -37.0498
26.053 | -35.8472 | -37.0553
26.0591 | -35.8476 | -37.0609
26.0652 | -35.848 | -37.0661
26.0713 | -35.8484 | -37.0715
26.0774 | -35.8488 | -37.0771
26.0835 | -35.8492 | -37.0829
26.0896 | -35.8497 | -37.0885
26.0957 | -35.8501 | -37.094
26.1018 | -35.8505 | -37.0994
26.1079 | -35.8509 | -37.1051
26.114 | -35.8513 | -37.1107
26.1201 | -35.8517 | -37.1164
26.1262 | -35.8521 | -37.1218
26.1323 | -35.8526 | -37.1272
26.1384 | -35.853 | -37.1325
26.1445 | -35.8534 | -37.1383
26.1507 | -35.8538 | -37.1439
26.1568 | -35.8542 | -37.1496
26.1629 | -35.8547 | -37.1553
26.169 | -35.8551 | -37.1608
26.1751 | -35.8555 | -37.1662
26.1812 | -35.8559 | -37.1718
26.1873 | -35.8564 | -37.1774
26.1934 | -35.8568 | -37.1828
26.1995 | -35.8572 | -37.1885
26.2056 | -35.8576 | -37.194
26.2117 | -35.858 | -37.1995
26.2178 | -35.8585 | -37.2051
26.2239 | -35.8589 | -37.2106
26.23 | -35.8593 | -37.2163
26.2361 | -35.8597 | -37.2219
26.2422 | -35.8601 | -37.2274
26.2483 | -35.8605 | -37.233
26.2544 | -35.8609 | -37.2385
26.2605 | -35.8613 | -37.2439
26.2666 | -35.8618 | -37.2493
26.2727 | -35.8622 | -37.2547
26.2788 | -35.8627 | -37.2603
26.2849 | -35.8631 | -37.2659
26.2911 | -35.8635 | -37.2714
26.2972 | -35.864 | -37.277
26.3033 | -35.8644 | -37.2826
26.3094 | -35.8648 | -37.2881
26.3155 | -35.8652 | -37.2933
26.3216 | -35.8656 | -37.2989
26.3277 | -35.866 | -37.3046
26.3338 | -35.8664 | -37.31
26.3399 | -35.8669 | -37.3155
26.346 | -35.8673 | -37.3206
26.3521 | -35.8678 | -37.3262
26.3582 | -35.8682 | -37.3316
26.3643 | -35.8686 | -37.3371
26.3704 | -35.869 | -37.3427
26.3765 | -35.8694 | -37.3483
26.3826 | -35.8698 | -37.3536
26.3887 | -35.8702 | -37.359
26.3948 | -35.8706 | -37.3646
26.4009 | -35.8711 | -37.3702
26.407 | -35.8715 | -37.3759
26.4131 | -35.8719 | -37.3816
26.4192 | -35.8723 | -37.387
26.4253 | -35.8727 | -37.3926
26.4314 | -35.8732 | -37.3981
26.4376 | -35.8736 | -37.4035
26.4437 | -35.874 | -37.4092
26.4498 | -35.8745 | -37.4148
26.4559 | -35.8749 | -37.4201
26.462 | -35.8753 | -37.4254
26.4681 | -35.8757 | -37.4309
26.4742 | -35.8761 | -37.4364
26.4803 | -35.8766 | -37.442
26.4864 | -35.877 | -37.4473
26.4925 | -35.8774 | -37.4528
26.4986 | -35.8778 | -37.4584
26.5047 | -35.8782 | -37.4639
26.5108 | -35.8786 | -37.4693
26.5169 | -35.879 | -37.475
26.523 | -35.8794 | -37.4802
26.5291 | -35.8798 | -37.4855
26.5352 | -35.8802 | -37.4911
26.5413 | -35.8806 | -37.4968
26.5474 | -35.8811 | -37.5025
26.5535 | -35.8815 | -37.508
26.5596 | -35.8819 | -37.5136
26.5657 | -35.8823 | -37.519
26.5718 | -35.8827 | -37.5245
26.578 | -35.8831 | -37.5303
26.5841 | -35.8835 | -37.5357
26.5902 | -35.8839 | -37.5411
26.5963 | -35.8843 | -37.5465
26.6024 | -35.8847 | -37.5523
26.6085 | -35.8851 | -37.5581
26.6146 | -35.8855 | -37.5636
26.6207 | -35.8859 | -37.5691
26.6268 | -35.8863 | -37.5745
26.6329 | -35.8867 | -37.5796
26.639 | -35.8871 | -37.5851
26.6451 | -35.8875 | -37.5906
26.6512 | -35.8879 | -37.5961
26.6573 | -35.8883 | -37.6015
26.6634 | -35.8888 | -37.6068
26.6695 | -35.8891 | -37.6125
26.6756 | -35.8896 | -37.6183
26.6817 | -35.89 | -37.6239
26.6878 | -35.8904 | -37.6295
26.6939 | -35.8908 | -37.6349
26.7 | -35.8911 | -37.6402
26.7061 | -35.8916 | -37.6457
26.7122 | -35.892 | -37.6513
26.7183 | -35.8924 | -37.6566
26.7245 | -35.8928 | -37.6621
26.7306 | -35.8932 | -37.6674
26.7367 | -35.8936 | -37.6726
26.7428 | -35.894 | -37.6782
26.7489 | -35.8944 | -37.6837
26.755 | -35.8948 | -37.6892
26.7611 | -35.8952 | -37.6949
26.7672 | -35.8956 | -37.7005
26.7733 | -35.8961 | -37.7059
26.7794 | -35.8965 | -37.7114
26.7855 | -35.8969 | -37.7167
26.7916 | -35.8973 | -37.7219
26.7977 | -35.8978 | -37.7272
26.8038 | -35.8981 | -37.7327
26.8099 | -35.8986 | -37.7382
26.816 | -35.8989 | -37.7437
26.8221 | -35.8994 | -37.7492
26.8282 | -35.8997 | -37.7548
26.8343 | -35.9001 | -37.7602
26.8404 | -35.9005 | -37.7654
26.8465 | -35.9009 | -37.7706
26.8526 | -35.9013 | -37.7763
26.8587 | -35.9017 | -37.7818
26.8649 | -35.9021 | -37.787
26.871 | -35.9025 | -37.7926
26.8771 | -35.9029 | -37.7975
26.8832 | -35.9033 | -37.8029
26.8893 | -35.9038 | -37.8084
26.8954 | -35.9042 | -37.8141
26.9015 | -35.9046 | -37.8195
26.9076 | -35.9049 | -37.8253
26.9137 | -35.9053 | -37.8306
26.9198 | -35.9057 | -37.8362
26.9259 | -35.9061 | -37.8418
26.932 | -35.9065 | -37.8472
26.9381 | -35.9069 | -37.8526
26.9442 | -35.9073 | -37.8577
26.9503 | -35.9077 | -37.863
26.9564 | -35.9081 | -37.8686
26.9625 | -35.9085 | -37.8741
26.9686 | -35.9089 | -37.8795
26.9747 | -35.9093 | -37.8849
26.9808 | -35.9097 | -37.8903
26.9869 | -35.9101 | -37.8957
26.993 | -35.9105 | -37.9012
26.9991 | -35.9108 | -37.9065
27.0052 | -35.9112 | -37.9121
27.0114 | -35.9116 | -37.9174
27.0175 | -35.912 | -37.9227
27.0236 | -35.9124 | -37.928
27.0297 | -35.9128 | -37.9334
27.0358 | -35.9132 | -37.9388
27.0419 | -35.9136 | -37.9443
27.048 | -35.914 | -37.9498
27.0541 | -35.9144 | -37.9551
27.0602 | -35.9148 | -37.9606
27.0663 | -35.9152 | -37.9661
27.0724 | -35.9156 | -37.9719
27.0785 | -35.916 | -37.9775
27.0846 | -35.9164 | -37.9829
27.0907 | -35.9168 | -37.9886
27.0968 | -35.9172 | -37.9943
27.1029 | -35.9175 | -38
27.109 | -35.9179 | -38.0055
27.1151 | -35.9183 | -38.0109
27.1212 | -35.9187 | -38.0161
27.1273 | -35.9191 | -38.0216
27.1334 | -35.9195 | -38.0271
27.1395 | -35.9199 | -38.0324
27.1456 | -35.9203 | -38.0378
27.1518 | -35.9207 | -38.0432
27.1579 | -35.9211 | -38.0484
27.164 | -35.9215 | -38.0535
27.1701 | -35.9219 | -38.0588
27.1762 | -35.9223 | -38.0643
27.1823 | -35.9227 | -38.0693
27.1884 | -35.9231 | -38.0748
27.1945 | -35.9235 | -38.0801
27.2006 | -35.9238 | -38.0854
27.2067 | -35.9242 | -38.0905
27.2128 | -35.9246 | -38.096
27.2189 | -35.925 | -38.1016
27.225 | -35.9254 | -38.1072
27.2311 | -35.9258 | -38.1125
27.2372 | -35.9262 | -38.1177
27.2433 | -35.9266 | -38.1228
27.2494 | -35.9269 | -38.1279
27.2555 | -35.9273 | -38.1333
27.2616 | -35.9277 | -38.1385
27.2677 | -35.9281 | -38.144
27.2738 | -35.9285 | -38.1494
27.2799 | -35.9289 | -38.1547
27.286 | -35.9293 | -38.1602
27.2921 | -35.9297 | -38.1657
27.2983 | -35.9301 | -38.1708
27.3044 | -35.9305 | -38.1764
27.3105 | -35.9309 | -38.1818
27.3166 | -35.9313 | -38.1873
27.3227 | -35.9317 | -38.1927
27.3288 | -35.9321 | -38.1979
27.3349 | -35.9325 | -38.2031
27.341 | -35.9329 | -38.2087
27.3471 | -35.9332 | -38.2144
27.3532 | -35.9336 | -38.2199
27.3593 | -35.9341 | -38.2257
27.3654 | -35.9344 | -38.2312
27.3715 | -35.9348 | -38.2364
27.3776 | -35.9352 | -38.242
27.3837 | -35.9356 | -38.2472
27.3898 | -35.936 | -38.2527
27.3959 | -35.9364 | -38.2579
27.402 | -35.9368 | -38.2629
27.4081 | -35.9372 | -38.268
27.4142 | -35.9376 | -38.2733
27.4203 | -35.938 | -38.2787
27.4264 | -35.9384 | -38.2841
27.4325 | -35.9388 | -38.2892
27.4387 | -35.9392 | -38.2947
27.4448 | -35.9396 | -38.3
27.4509 | -35.94 | -38.3053
27.457 | -35.9404 | -38.3106
27.4631 | -35.9407 | -38.3161
27.4692 | -35.9411 | -38.3214
27.4753 | -35.9415 | -38.3264
27.4814 | -35.9419 | -38.3317
27.4875 | -35.9424 | -38.3372
27.4936 | -35.9428 | -38.3424
27.4997 | -35.9432 | -38.3479
27.5058 | -35.9436 | -38.3535
27.5119 | -35.944 | -38.3587
27.518 | -35.9445 | -38.3642
27.5241 | -35.9448 | -38.3698
27.5302 | -35.9452 | -38.3751
27.5363 | -35.9456 | -38.3806
27.5424 | -35.9459 | -38.3852
27.5485 | -35.9464 | -38.3906
27.5546 | -35.9468 | -38.396
27.5607 | -35.9471 | -38.4015
27.5668 | -35.9476 | -38.4067
27.5729 | -35.948 | -38.4121
27.5791 | -35.9483 | -38.4173
27.5852 | -35.9487 | -38.4225
27.5913 | -35.9491 | -38.4278
27.5974 | -35.9495 | -38.433
27.6035 | -35.9499 | -38.4384
27.6096 | -35.9502 | -38.4436
27.6157 | -35.9506 | -38.449
27.6218 | -35.951 | -38.4547
27.6279 | -35.9514 | -38.4603
27.634 | -35.9518 | -38.4655
27.6401 | -35.9522 | -38.4708
27.6462 | -35.9527 | -38.4763
27.6523 | -35.9531 | -38.481
27.6584 | -35.9535 | -38.4864
27.6645 | -35.9538 | -38.4916
27.6706 | -35.9542 | -38.4969
27.6767 | -35.9545 | -38.5019
27.6828 | -35.9549 | -38.5073
27.6889 | -35.9553 | -38.5128
27.695 | -35.9557 | -38.5182
27.7011 | -35.9561 | -38.5235
27.7072 | -35.9565 | -38.529
27.7133 | -35.9569 | -38.5343
27.7194 | -35.9573 | -38.54
27.7256 | -35.9577 | -38.5455
27.7317 | -35.9582 | -38.5509
27.7378 | -35.9586 | -38.5563
27.7439 | -35.959 | -38.5618
27.75 | -35.9594 | -38.5673
27.7561 | -35.9598 | -38.5724
27.7622 | -35.9602 | -38.578
27.7683 | -35.9606 | -38.5829
27.7744 | -35.961 | -38.5881
27.7805 | -35.9613 | -38.5935
27.7866 | -35.9617 | -38.5991
27.7927 | -35.9621 | -38.604
27.7988 | -35.9624 | -38.6095
27.8049 | -35.9629 | -38.615
27.811 | -35.9632 | -38.6205
27.8171 | -35.9636 | -38.6262
27.8232 | -35.9641 | -38.6312
27.8293 | -35.9644 | -38.6366
27.8354 | -35.9648 | -38.6419
27.8415 | -35.9652 | -38.6468
27.8476 | -35.9656 | -38.6523
27.8537 | -35.966 | -38.6575
27.8598 | -35.9664 | -38.6634
27.866 | -35.9668 | -38.6687
27.8721 | -35.9672 | -38.6742
27.8782 | -35.9676 | -38.6794
27.8843 | -35.968 | -38.6845
27.8904 | -35.9684 | -38.6898
27.8965 | -35.9688 | -38.6951
27.9026 | -35.9692 | -38.7004
27.9087 | -35.9696 | -38.7054
27.9148 | -35.97 | -38.7108
27.9209 | -35.9703 | -38.7158
27.927 | -35.9708 | -38.7213
27.9331 | -35.9712 | -38.7271
27.9392 | -35.9716 | -38.7322
27.9453 | -35.972 | -38.7378
27.9514 | -35.9724 | -38.743
27.9575 | -35.9727 | -38.7485
27.9636 | -35.9731 | -38.7537
27.9697 | -35.9735 | -38.7594
27.9758 | -35.9739 | -38.7647
27.9819 | -35.9743 | -38.7698
27.988 | -35.9747 | -38.7752
27.9941 | -35.9751 | -38.7806
28.0002 | -35.9755 | -38.7857
28.0063 | -35.9758 | -38.791
28.0125 | -35.9762 | -38.7964
28.0186 | -35.9766 | -38.8014
28.0247 | -35.9771 | -38.8065
28.0308 | -35.9775 | -38.8116
28.0369 | -35.9778 | -38.8167
28.043 | -35.9782 | -38.8219
28.0491 | -35.9786 | -38.8273
28.0552 | -35.979 | -38.8326
28.0613 | -35.9794 | -38.8373
28.0674 | -35.9798 | -38.8426
28.0735 | -35.9802 | -38.848

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.8256 | -43.439
21.4908 | -35.7548 | -35.7951
16.4948 | -35.0858 | -28.2812
12.6602 | -34.6908 | -21.8939
9.7171 | -34.454 | -16.8851
7.4582 | -34.2938 | -12.8349
5.7244 | -34.1828 | -9.5491
4.3936 | -34.1065 | -6.8947
3.3722 | -34.0901 | -4.8069
2.5883 | -34.0858 | -3.2214
1.9866 | -34.0845 | -1.955
1.5248 | -34.0962 | -0.93631
1.1703 | -34.0953 | -0.27419
1 | -34.097 | 0.058074
0.89822 | -34.0941 | 0.33194
0.68941 | -34.0816 | 0.92154
0.52919 | -34.0841 | 1.3794
0.40612 | -34.0845 | 1.9936
0.31171 | -34.0826 | 2.4494
0.23928 | -34.1005 | 3.1679
0.18358 | -34.0884 | 3.5912
0.14091 | -34.0966 | 4.465
0.10819 | -34.1206 | 5.6216
0.083017 | -34.1511 | 7.0925
0.063705 | -34.188 | 8.9632
0.048923 | -34.2443 | 11.4493
0.037479 | -34.3588 | 14.6429
0.028801 | -34.5274 | 18.669
0.022078 | -34.8407 | 24.1217
0.017023 | -35.3303 | 30.5655
0.013018 | -36.1342 | 38.2403
0.010014 | -37.2554 | 46.1766

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003663 | -11.9718
0.007326 | -31.5201
0.010989 | -35.5569
0.014651 | -39.6799
0.018314 | -41.7776
0.021977 | -43.5723
0.02564 | -46.0627
0.029303 | -46.1743
0.032966 | -47.635
0.036628 | -49.9178
0.040291 | -51.974
0.043954 | -53.8863
0.047617 | -54.4776
0.05128 | -54.7498
0.054943 | -55.3661
0.058605 | -56.2159
0.062268 | -57.3165
0.065931 | -58.2195
0.069594 | -58.0749
0.073257 | -58.4536
0.07692 | -59.4228
0.080582 | -59.7842
0.084245 | -61.0561
0.087908 | -61.8066
0.091571 | -61.7524
0.095234 | -61.5665
0.098897 | -61.4538
0.10256 | -61.3247
0.10622 | -61.7255
0.10988 | -62.582
0.11355 | -63.4652
0.11721 | -64.0227
0.12087 | -64.8382
0.12454 | -64.5096
0.1282 | -64.7974
0.13186 | -65.0638
0.13553 | -65.9635
0.13919 | -66.158
0.14285 | -66.1085
0.14651 | -67.208
0.15018 | -67.1859
0.15384 | -66.6988
0.1575 | -66.5729
0.16117 | -67.5735
0.16483 | -67.6814
0.16849 | -67.5208
0.17215 | -67.8556
0.17582 | -67.3561
0.17948 | -68.537
0.18314 | -69.8001
0.1868 | -70.1315
0.19047 | -69.7504
0.19413 | -69.8317
0.19779 | -70.2645
0.20146 | -70.8871
0.20512 | -70.8812
0.20878 | -69.8712
0.21244 | -70.1108
0.21611 | -70.8037
0.21977 | -71.0591
0.22343 | -72.4371
0.2271 | -72.7235
0.23076 | -72.1194
0.23442 | -72.0461
0.23808 | -71.8338
0.24175 | -71.5295
0.24541 | -72.7784
0.24907 | -72.7479
0.25274 | -73.1933
0.2564 | -73.1934
0.26006 | -73.5337
0.26372 | -73.4874
0.26739 | -73.154
0.27105 | -73.3788
0.27471 | -73.7166
0.27838 | -73.5528
0.28204 | -73.6727
0.2857 | -73.583
0.28936 | -73.1499
0.29303 | -74.0272
0.29669 | -74.141
0.30035 | -74.7182
0.30401 | -74.8115
0.30768 | -74.4105
0.31134 | -74.9635
0.315 | -75.1084
0.31867 | -74.5281
0.32233 | -75.3689
0.32599 | -75.3478
0.32965 | -76.0046
0.33332 | -74.6592
0.33698 | -74.6593
0.34064 | -77.0731
0.34431 | -76.1972
0.34797 | -76.1869
0.35163 | -76.2611
0.3553 | -76.397
0.35896 | -76.0983
0.36262 | -76.0198
0.36628 | -76.1468
0.36995 | -75.6627
0.37361 | -76.408
0.37727 | -77.7244
0.38094 | -77.4005
0.3846 | -77.2237
0.38826 | -77.3257
0.39192 | -77.8002
0.39559 | -77.41
0.39925 | -77.0465
0.40291 | -76.2448
0.40658 | -76.9792
0.41024 | -77.6178
0.4139 | -77.3159
0.41756 | -77.2517
0.42123 | -78.199
0.42489 | -79.1017
0.42855 | -78.4775
0.43222 | -77.0003
0.43588 | -77.2822
0.43954 | -76.6709
0.4432 | -78.042
0.44687 | -78.8386
0.45053 | -77.6339
0.45419 | -77.7727
0.45785 | -78.5926
0.46152 | -79.7382
0.46518 | -79.5018
0.46884 | -78.9658
0.47251 | -77.5461
0.47617 | -77.6435
0.47983 | -78.7641
0.48349 | -79.5558
0.48716 | -79.4698
0.49082 | -79.077
0.49448 | -78.6461
0.49815 | -79.6475
0.50181 | -80.2021
0.50547 | -79.825
0.50913 | -79.2335
0.5128 | -79.3829
0.51646 | -79.8104
0.52012 | -79.6869
0.52379 | -80.6911
0.52745 | -80.9388
0.53111 | -81.416
0.53477 | -81.4319
0.53844 | -80.8095
0.5421 | -80.5273
0.54576 | -80.3792
0.54943 | -80.2582
0.55309 | -80.4329
0.55675 | -80.4179
0.56041 | -80.3808
0.56408 | -80.7788
0.56774 | -81.3246
0.5714 | -81.188
0.57507 | -80.8419
0.57873 | -81.0605
0.58239 | -82.0001
0.58605 | -82.943
0.58972 | -82.7429
0.59338 | -82.2785
0.59704 | -81.9557
0.60071 | -81.5875
0.60437 | -80.9784
0.60803 | -80.9572
0.61169 | -81.9634
0.61536 | -82.2991
0.61902 | -81.9689
0.62268 | -81.7233
0.62635 | -81.6478
0.63001 | -81.6833
0.63367 | -81.6772
0.63733 | -81.8297
0.641 | -82.0132
0.64466 | -82.1663
0.64832 | -82.1351
0.65199 | -81.8153
0.65565 | -81.7406
0.65931 | -82.3164
0.66297 | -82.8646
0.66664 | -82.6669
0.6703 | -82.175
0.67396 | -82.1925
0.67763 | -82.8151
0.68129 | -83.0594
0.68495 | -83.2932
0.68861 | -83.2567
0.69228 | -82.8054
0.69594 | -82.8471
0.6996 | -83.4469
0.70327 | -83.9883
0.70693 | -84.0843
0.71059 | -84.0364
0.71425 | -83.8065
0.71792 | -83.4514
0.72158 | -82.9763
0.72524 | -82.3835
0.7289 | -82.058
0.73257 | -82.4345
0.73623 | -82.954
0.73989 | -83.2178
0.74356 | -83.8954
0.74722 | -84.4955
0.75088 | -84.4708
0.75454 | -84.5229
0.75821 | -84.9746
0.76187 | -85.052
0.76553 | -84.2088
0.7692 | -83.4228
0.77286 | -83.3327
0.77652 | -83.4825
0.78018 | -83.458
0.78385 | -83.2531
0.78751 | -83.3475
0.79117 | -83.3524
0.79483 | -83.1678
0.7985 | -83.46
0.80216 | -84.0387
0.80582 | -84.2838
0.80949 | -84.2408
0.81315 | -84.6545
0.81681 | -84.8128
0.82047 | -84.392
0.82414 | -83.8368
0.8278 | -83.5856
0.83146 | -83.8836
0.83513 | -84.2352
0.83879 | -84.3531
0.84245 | -84.3983
0.84611 | -84.6896
0.84978 | -84.8544
0.85344 | -84.7958
0.8571 | -84.7374
0.86077 | -84.6659
0.86443 | -85.0811
0.86809 | -85.5066
0.87175 | -85.0255
0.87542 | -84.3937
0.87908 | -84.2627
0.88274 | -84.4986
0.88641 | -84.9792
0.89007 | -85.6778
0.89373 | -85.7755
0.8974 | -85.5457
0.90106 | -85.1624
0.90472 | -85.078
0.90838 | -85.3608
0.91205 | -85.7218
0.91571 | -85.7005
0.91937 | -85.2085
0.92304 | -84.8654
0.9267 | -84.9258
0.93036 | -85.1906
0.93402 | -85.253
0.93769 | -85.2101
0.94135 | -85.2742
0.94501 | -85.1167
0.94868 | -85.1371
0.95234 | -85.4591
0.956 | -85.9358
0.95966 | -86.2824
0.96333 | -86.0287
0.96699 | -85.9705
0.97065 | -85.9852
0.97431 | -85.8025
0.97798 | -85.7429
0.98164 | -85.5848
0.9853 | -85.3642
0.98897 | -85.376
0.99263 | -85.572
0.99629 | -85.9934
0.99995 | -86.2091
1.0036 | -86.2795
1.0073 | -86.4518
1.0109 | -86.8315
1.0146 | -87.002
1.0183 | -87.0242
1.0219 | -86.7497
1.0256 | -86.0911
1.0293 | -86.0276
1.0329 | -86.4374
1.0366 | -86.5867
1.0402 | -86.3529
1.0439 | -86.1241
1.0476 | -86.4616
1.0512 | -87.3174
1.0549 | -87.1855
1.0586 | -87.0329
1.0622 | -86.9552
1.0659 | -87.0599
1.0695 | -87.2277
1.0732 | -87.3538
1.0769 | -87.425
1.0805 | -87.5218
1.0842 | -87.6265
1.0879 | -87.6478
1.0915 | -87.6042
1.0952 | -87.5912
1.0989 | -87.6728
1.1025 | -87.8468
1.1062 | -87.9707
1.1098 | -87.9786
1.1135 | -88.0233
1.1172 | -88.1379
1.1208 | -88.1375
1.1245 | -88.0134
1.1282 | -87.9032
1.1318 | -87.8344
1.1355 | -87.7638
1.1391 | -87.711
1.1428 | -87.6885
1.1465 | -87.6189
1.1501 | -87.4338
1.1538 | -87.1373
1.1575 | -86.8388
1.1611 | -86.7443
1.1648 | -86.9419
1.1684 | -87.254
1.1721 | -87.4487
1.1758 | -87.5745
1.1794 | -87.6441
1.1831 | -87.5867
1.1868 | -87.6198
1.1904 | -87.8989
1.1941 | -88.1963
1.1977 | -88.2449
1.2014 | -88.08
1.2051 | -87.9598
1.2087 | -88.0894
1.2124 | -88.4284
1.2161 | -88.6483
1.2197 | -88.575
1.2234 | -88.3461
1.227 | -88.1287
1.2307 | -88.1204
1.2344 | -88.3769
1.238 | -88.614
1.2417 | -88.6413
1.2454 | -88.6405
1.249 | -88.8387
1.2527 | -89.2691
1.2564 | -89.6872
1.26 | -89.6355
1.2637 | -89.1237
1.2673 | -88.698
1.271 | -88.7266
1.2747 | -89.0835
1.2783 | -89.4044
1.282 | -89.5065
1.2857 | -89.3713
1.2893 | -89.0432
1.293 | -88.7349
1.2966 | -88.6286
1.3003 | -88.7511
1.304 | -88.9829
1.3076 | -89.2286
1.3113 | -89.4652
1.315 | -89.5353
1.3186 | -89.2816
1.3223 | -88.8491
1.3259 | -88.5712
1.3296 | -88.6599
1.3333 | -89.0068
1.3369 | -89.2784
1.3406 | -89.2853
1.3443 | -89.0641
1.3479 | -88.7165
1.3516 | -88.5072
1.3552 | -88.6604
1.3589 | -89.0045
1.3626 | -89.1764
1.3662 | -89.2654
1.3699 | -89.599
1.3736 | -90.1147
1.3772 | -90.4544
1.3809 | -90.4347
1.3846 | -90.147
1.3882 | -89.8027
1.3919 | -89.5261
1.3955 | -89.339
1.3992 | -89.3137
1.4029 | -89.5564
1.4065 | -89.9448
1.4102 | -90.1366
1.4139 | -90.1069
1.4175 | -90.0639
1.4212 | -89.9352
1.4248 | -89.7305
1.4285 | -89.6855
1.4322 | -89.8785
1.4358 | -90.202
1.4395 | -90.4615
1.4432 | -90.569
1.4468 | -90.6501
1.4505 | -90.7401
1.4541 | -90.7539
1.4578 | -90.6387
1.4615 | -90.4086
1.4651 | -90.1162
1.4688 | -89.9027
1.4725 | -89.8019
1.4761 | -89.7024
1.4798 | -89.6659
1.4834 | -89.7825
1.4871 | -89.9476
1.4908 | -90.0739
1.4944 | -90.0852
1.4981 | -89.951
1.5018 | -89.807
1.5054 | -89.7685
1.5091 | -89.8262
1.5128 | -89.9216
1.5164 | -89.9809
1.5201 | -90.0294
1.5237 | -90.1895
1.5274 | -90.3875
1.5311 | -90.4837
1.5347 | -90.5674
1.5384 | -90.7229
1.5421 | -90.8005
1.5457 | -90.665
1.5494 | -90.4402
1.553 | -90.3275
1.5567 | -90.3773
1.5604 | -90.4572
1.564 | -90.3714
1.5677 | -90.2185
1.5714 | -90.2533
1.575 | -90.4842
1.5787 | -90.7922
1.5823 | -91.0484
1.586 | -91.2001
1.5897 | -91.1427
1.5933 | -90.9302
1.597 | -90.7101
1.6007 | -90.5899
1.6043 | -90.6084
1.608 | -90.7406
1.6116 | -90.8983
1.6153 | -90.9539
1.619 | -90.8588
1.6226 | -90.7584
1.6263 | -90.7954
1.63 | -90.9917
1.6336 | -91.2291
1.6373 | -91.3592
1.641 | -91.352
1.6446 | -91.2737
1.6483 | -91.2392
1.6519 | -91.3135
1.6556 | -91.4647
1.6593 | -91.64
1.6629 | -91.7585
1.6666 | -91.7654
1.6703 | -91.69
1.6739 | -91.5633
1.6776 | -91.3963
1.6812 | -91.1647
1.6849 | -90.8845
1.6886 | -90.6381
1.6922 | -90.58
1.6959 | -90.8171
1.6996 | -91.2854
1.7032 | -91.7182
1.7069 | -91.8819
1.7105 | -91.7756
1.7142 | -91.5808
1.7179 | -91.4571
1.7215 | -91.4714
1.7252 | -91.6127
1.7289 | -91.8248
1.7325 | -91.9842
1.7362 | -91.9781
1.7398 | -91.8097
1.7435 | -91.5768
1.7472 | -91.4082
1.7508 | -91.4007
1.7545 | -91.5358
1.7582 | -91.7438
1.7618 | -91.9554
1.7655 | -92.1022
1.7692 | -92.099
1.7728 | -91.967
1.7765 | -91.9022
1.7801 | -92.0371
1.7838 | -92.3199
1.7875 | -92.5826
1.7911 | -92.6659
1.7948 | -92.55
1.7985 | -92.3512
1.8021 | -92.1884
1.8058 | -92.1077
1.8094 | -92.0699
1.8131 | -92.026
1.8168 | -91.9883
1.8204 | -91.9892
1.8241 | -91.9964
1.8278 | -91.927
1.8314 | -91.7915
1.8351 | -91.7306
1.8387 | -91.8276
1.8424 | -92.0531
1.8461 | -92.3151
1.8497 | -92.4922
1.8534 | -92.4929
1.8571 | -92.3112
1.8607 | -92.0798
1.8644 | -91.977
1.868 | -92.056
1.8717 | -92.255
1.8754 | -92.4643
1.879 | -92.5714
1.8827 | -92.5188
1.8864 | -92.3655
1.89 | -92.2582
1.8937 | -92.2489
1.8973 | -92.2458
1.901 | -92.2205
1.9047 | -92.1896
1.9083 | -92.1691
1.912 | -92.1198
1.9157 | -92.0294
1.9193 | -91.9535
1.923 | -91.9585
1.9267 | -92.0495
1.9303 | -92.2308
1.934 | -92.483
1.9376 | -92.7208
1.9413 | -92.7937
1.945 | -92.6846
1.9486 | -92.5356
1.9523 | -92.4799
1.956 | -92.5274
1.9596 | -92.5784
1.9633 | -92.5667
1.9669 | -92.5588
1.9706 | -92.6776
1.9743 | -92.9344
1.9779 | -93.1995
1.9816 | -93.3142
1.9853 | -93.2743
1.9889 | -93.2138
1.9926 | -93.207
1.9962 | -93.152
1.9999 | -92.9547
2.0036 | -92.6481
2.0072 | -92.3894
2.0109 | -92.3364
2.0146 | -92.521
2.0182 | -92.7819
2.0219 | -92.8506
2.0255 | -92.6072
2.0292 | -92.228
2.0329 | -91.9628
2.0365 | -91.9606
2.0402 | -92.2282
2.0439 | -92.6946
2.0475 | -93.2577
2.0512 | -93.732
2.0549 | -93.8712
2.0585 | -93.6014
2.0622 | -93.1043
2.0658 | -92.6565
2.0695 | -92.4395
2.0732 | -92.5187
2.0768 | -92.7811
2.0805 | -92.9929
2.0842 | -93.0543
2.0878 | -93.0903
2.0915 | -93.2256
2.0951 | -93.4378
2.0988 | -93.5922
2.1025 | -93.6026
2.1061 | -93.4952
2.1098 | -93.3586
2.1135 | -93.3746
2.1171 | -93.3955
2.1208 | -93.4438
2.1244 | -93.4602
2.1281 | -93.3986
2.1318 | -93.2658
2.1354 | -93.1325
2.1391 | -93.0549
2.1428 | -93.0478
2.1464 | -93.1262
2.1501 | -93.2552
2.1537 | -93.3861
2.1574 | -93.4871
2.1611 | -93.5222
2.1647 | -93.4806
2.1684 | -93.3565
2.1721 | -93.1928
2.1757 | -93.0676
2.1794 | -93.0571
2.1831 | -93.1753
2.1867 | -93.3633
2.1904 | -93.5334
2.194 | -93.6161
2.1977 | -93.5982
2.2014 | -93.548
2.205 | -93.5168
2.2087 | -93.5281
2.2124 | -93.5598
2.216 | -93.6181
2.2197 | -93.732
2.2233 | -93.9427
2.227 | -94.2153
2.2307 | -94.4535
2.2343 | -94.56
2.238 | -94.46
2.2417 | -94.1776
2.2453 | -93.7854
2.249 | -93.4072
2.2526 | -93.1454
2.2563 | -93.056
2.26 | -93.1329
2.2636 | -93.3157
2.2673 | -93.5255
2.271 | -93.7016
2.2746 | -93.8098
2.2783 | -93.8614
2.2819 | -93.8788
2.2856 | -93.8979
2.2893 | -93.9412
2.2929 | -94.0288
2.2966 | -94.1459
2.3003 | -94.2508
2.3039 | -94.2864
2.3076 | -94.208
2.3112 | -94.0132
2.3149 | -93.765
2.3186 | -93.5426
2.3222 | -93.3941
2.3259 | -93.342
2.3296 | -93.3748
2.3332 | -93.4725
2.3369 | -93.6156
2.3406 | -93.7764
2.3442 | -93.9103
2.3479 | -93.998
2.3515 | -94.0508
2.3552 | -94.1065
2.3589 | -94.179
2.3625 | -94.251
2.3662 | -94.2836
2.3699 | -94.2685
2.3735 | -94.2233
2.3772 | -94.1768
2.3808 | -94.1556
2.3845 | -94.1491
2.3882 | -94.143
2.3918 | -94.1398
2.3955 | -94.1512
2.3992 | -94.1823
2.4028 | -94.1987
2.4065 | -94.1732
2.4101 | -94.0685
2.4138 | -93.9078
2.4175 | -93.7771
2.4211 | -93.7478
2.4248 | -93.8498
2.4285 | -94.036
2.4321 | -94.2336
2.4358 | -94.3642
2.4394 | -94.3943
2.4431 | -94.3398
2.4468 | -94.2583
2.4504 | -94.2207
2.4541 | -94.2451
2.4578 | -94.3035
2.4614 | -94.3614
2.4651 | -94.3523
2.4688 | -94.237
2.4724 | -94.0487
2.4761 | -93.8761
2.4797 | -93.7712
2.4834 | -93.7464
2.4871 | -93.754
2.4907 | -93.7682
2.4944 | -93.7957
2.4981 | -93.8403
2.5017 | -93.9071
2.5054 | -93.9904
2.509 | -94.0725
2.5127 | -94.1328
2.5164 | -94.1402
2.52 | -94.099
2.5237 | -94.037
2.5274 | -93.9975
2.531 | -94.0159
2.5347 | -94.0875
2.5383 | -94.2133
2.542 | -94.3668
2.5457 | -94.5204
2.5493 | -94.6603
2.553 | -94.7799
2.5567 | -94.8984
2.5603 | -95.0119
2.564 | -95.0844
2.5676 | -95.0949
2.5713 | -95.0106
2.575 | -94.8398
2.5786 | -94.6481
2.5823 | -94.5295
2.586 | -94.5486
2.5896 | -94.6782
2.5933 | -94.8475
2.597 | -94.9684
2.6006 | -95.0184
2.6043 | -95.032
2.6079 | -95.0172
2.6116 | -94.9966
2.6153 | -94.992
2.6189 | -94.9915
2.6226 | -94.9968
2.6263 | -94.9913
2.6299 | -94.9738
2.6336 | -94.9489
2.6372 | -94.933
2.6409 | -94.8722
2.6446 | -94.807
2.6482 | -94.7333
2.6519 | -94.6746
2.6556 | -94.6441
2.6592 | -94.6416
2.6629 | -94.6526
2.6665 | -94.6556
2.6702 | -94.6573
2.6739 | -94.6625
2.6775 | -94.6939
2.6812 | -94.7669
2.6849 | -94.9086
2.6885 | -95.1173
2.6922 | -95.3655
2.6958 | -95.6015
2.6995 | -95.7773
2.7032 | -95.8314
2.7068 | -95.7754
2.7105 | -95.6284
2.7142 | -95.4323
2.7178 | -95.2227
2.7215 | -95.0456
2.7252 | -94.9306
2.7288 | -94.9028
2.7325 | -94.9438
2.7361 | -95.0336
2.7398 | -95.1358
2.7435 | -95.2069
2.7471 | -95.2363
2.7508 | -95.2336
2.7545 | -95.198
2.7581 | -95.1564
2.7618 | -95.1114
2.7654 | -95.0691
2.7691 | -95.0389
2.7728 | -95.0309
2.7764 | -95.0545
2.7801 | -95.1225
2.7838 | -95.2023
2.7874 | -95.2487
2.7911 | -95.2384
2.7947 | -95.1421
2.7984 | -94.9851
2.8021 | -94.8211
2.8057 | -94.7061
2.8094 | -94.6748
2.8131 | -94.7233
2.8167 | -94.8358
2.8204 | -94.9781
2.824 | -95.1251
2.8277 | -95.2553
2.8314 | -95.3543
2.835 | -95.4372
2.8387 | -95.5223
2.8424 | -95.6172
2.846 | -95.6981
2.8497 | -95.7309
2.8533 | -95.684
2.857 | -95.5549
2.8607 | -95.3899
2.8643 | -95.2381
2.868 | -95.1431
2.8717 | -95.1156
2.8753 | -95.1366
2.879 | -95.179
2.8827 | -95.2041
2.8863 | -95.2011
2.89 | -95.1834
2.8936 | -95.173
2.8973 | -95.1902
2.901 | -95.2361
2.9046 | -95.312
2.9083 | -95.4092
2.912 | -95.5336
2.9156 | -95.6903
2.9193 | -95.8751
2.9229 | -96.0566
2.9266 | -96.1792
2.9303 | -96.1972
2.9339 | -96.1023
2.9376 | -95.9275
2.9413 | -95.719
2.9449 | -95.5423
2.9486 | -95.4617
2.9522 | -95.4842
2.9559 | -95.5867
2.9596 | -95.7449
2.9632 | -95.9023
2.9669 | -96.001
2.9706 | -96.0255
2.9742 | -95.9957
2.9779 | -95.9533
2.9815 | -95.9423
2.9852 | -95.9774
2.9889 | -96.0544
2.9925 | -96.1298
2.9962 | -96.1631
2.9999 | -96.1408
3.0035 | -96.0743
3.0072 | -95.9814
3.0109 | -95.9231
3.0145 | -95.9159
3.0182 | -95.9676
3.0218 | -96.0557
3.0255 | -96.1514
3.0292 | -96.22
3.0328 | -96.2366
3.0365 | -96.1867
3.0402 | -96.0818
3.0438 | -95.9471
3.0475 | -95.8137
3.0511 | -95.7185
3.0548 | -95.6805
3.0585 | -95.6983
3.0621 | -95.7547
3.0658 | -95.7883
3.0695 | -95.7693
3.0731 | -95.6969
3.0768 | -95.6015
3.0804 | -95.5166
3.0841 | -95.4639
3.0878 | -95.4422
3.0914 | -95.4365
3.0951 | -95.4265
3.0988 | -95.3918
3.1024 | -95.3495
3.1061 | -95.3014
3.1097 | -95.2631
3.1134 | -95.2392
3.1171 | -95.2369
3.1207 | -95.2754
3.1244 | -95.3803
3.1281 | -95.5456
3.1317 | -95.7342
3.1354 | -95.9104
3.1391 | -96.0373
3.1427 | -96.111
3.1464 | -96.1507
3.15 | -96.1731
3.1537 | -96.213
3.1574 | -96.2815
3.161 | -96.393
3.1647 | -96.5221
3.1684 | -96.6249
3.172 | -96.7201
3.1757 | -96.775
3.1793 | -96.7918
3.183 | -96.775
3.1867 | -96.7147
3.1903 | -96.6272
3.194 | -96.5194
3.1977 | -96.4147
3.2013 | -96.3159
3.205 | -96.2377
3.2086 | -96.1851
3.2123 | -96.1593
3.216 | -96.1691
3.2196 | -96.2162
3.2233 | -96.3006
3.227 | -96.4208
3.2306 | -96.552
3.2343 | -96.6719
3.2379 | -96.7523
3.2416 | -96.7804
3.2453 | -96.772
3.2489 | -96.733
3.2526 | -96.6656
3.2563 | -96.581
3.2599 | -96.4677
3.2636 | -96.3515
3.2673 | -96.2344
3.2709 | -96.12
3.2746 | -96.0199
3.2782 | -95.9439
3.2819 | -95.9022
3.2856 | -95.9103
3.2892 | -95.9549
3.2929 | -96.042
3.2966 | -96.1679
3.3002 | -96.315
3.3039 | -96.4871
3.3075 | -96.6708
3.3112 | -96.8506
3.3149 | -97.0036
3.3185 | -97.0894
3.3222 | -97.1171
3.3259 | -97.093
3.3295 | -97.0264
3.3332 | -96.964
3.3368 | -96.935
3.3405 | -96.9371
3.3442 | -96.965
3.3478 | -96.9931
3.3515 | -96.996
3.3552 | -96.9586
3.3588 | -96.8843
3.3625 | -96.7905
3.3661 | -96.6942
3.3698 | -96.604
3.3735 | -96.5397
3.3771 | -96.5083
3.3808 | -96.5205
3.3845 | -96.5738
3.3881 | -96.639
3.3918 | -96.7045
3.3954 | -96.7662
3.3991 | -96.8169
3.4028 | -96.8504
3.4064 | -96.8818
3.4101 | -96.9212
3.4138 | -96.9582
3.4174 | -96.9788
3.4211 | -96.9776
3.4248 | -96.9686
3.4284 | -96.9714
3.4321 | -96.995
3.4357 | -97.0379
3.4394 | -97.1023
3.4431 | -97.1699
3.4467 | -97.2088
3.4504 | -97.2165
3.4541 | -97.1929
3.4577 | -97.1639
3.4614 | -97.1337
3.465 | -97.1243
3.4687 | -97.1404
3.4724 | -97.1724
3.476 | -97.2167
3.4797 | -97.2631
3.4834 | -97.3099
3.487 | -97.3688
3.4907 | -97.4391
3.4943 | -97.4942
3.498 | -97.5392
3.5017 | -97.5769
3.5053 | -97.6033
3.509 | -97.624
3.5127 | -97.6268
3.5163 | -97.6149
3.52 | -97.5765
3.5236 | -97.5193
3.5273 | -97.4453
3.531 | -97.3568
3.5346 | -97.2801
3.5383 | -97.2301
3.542 | -97.2153
3.5456 | -97.2407
3.5493 | -97.2994
3.553 | -97.3606
3.5566 | -97.4106
3.5603 | -97.4373
3.5639 | -97.4455
3.5676 | -97.4455
3.5713 | -97.4409
3.5749 | -97.4379
3.5786 | -97.4446
3.5823 | -97.4585
3.5859 | -97.4593
3.5896 | -97.436
3.5932 | -97.3982
3.5969 | -97.3587
3.6006 | -97.3369
3.6042 | -97.3437
3.6079 | -97.3831
3.6116 | -97.4291
3.6152 | -97.4686
3.6189 | -97.5022
3.6225 | -97.5367
3.6262 | -97.5883
3.6299 | -97.6641
3.6335 | -97.7561
3.6372 | -97.8418
3.6409 | -97.9116
3.6445 | -97.9516
3.6482 | -97.9476
3.6518 | -97.9047
3.6555 | -97.824
3.6592 | -97.7316
3.6628 | -97.6434
3.6665 | -97.5897
3.6702 | -97.5664
3.6738 | -97.5721
3.6775 | -97.5852
3.6812 | -97.5784
3.6848 | -97.5415
3.6885 | -97.4858
3.6921 | -97.4249
3.6958 | -97.4064
3.6995 | -97.3918
3.7031 | -97.4011
3.7068 | -97.4388
3.7105 | -97.5013
3.7141 | -97.584
3.7178 | -97.6771
3.7214 | -97.7645
3.7251 | -97.846
3.7288 | -97.9283
3.7324 | -97.9942
3.7361 | -98.051
3.7398 | -98.0876
3.7434 | -98.1069
3.7471 | -98.0976
3.7507 | -98.0604
3.7544 | -98.0041
3.7581 | -97.9528
3.7617 | -97.9087
3.7654 | -97.8577
3.7691 | -97.8065
3.7727 | -97.7596
3.7764 | -97.719
3.78 | -97.6805
3.7837 | -97.6663
3.7874 | -97.6703
3.791 | -97.6894
3.7947 | -97.7258
3.7984 | -97.7685
3.802 | -97.8201
3.8057 | -97.8639
3.8093 | -97.906
3.813 | -97.9491
3.8167 | -97.9816
3.8203 | -97.9965
3.824 | -98.0063
3.8277 | -98.0066
3.8313 | -97.9929
3.835 | -97.9523
3.8387 | -97.8952
3.8423 | -97.8231
3.846 | -97.7239
3.8496 | -97.6169
3.8533 | -97.5226
3.857 | -97.4602
3.8606 | -97.4465
3.8643 | -97.4871
3.868 | -97.5682
3.8716 | -97.6733
3.8753 | -97.7709
3.8789 | -97.8439
3.8826 | -97.8813
3.8863 | -97.8818
3.8899 | -97.8625
3.8936 | -97.8365
3.8973 | -97.8166
3.9009 | -97.8059
3.9046 | -97.8132
3.9082 | -97.8463
3.9119 | -97.897
3.9156 | -97.9625
3.9192 | -98.0177
3.9229 | -98.063
3.9266 | -98.0889
3.9302 | -98.1081
3.9339 | -98.1165
3.9375 | -98.1337
3.9412 | -98.1717
3.9449 | -98.2301
3.9485 | -98.3054
3.9522 | -98.3938
3.9559 | -98.4775
3.9595 | -98.5547
3.9632 | -98.6198
3.9669 | -98.6586
3.9705 | -98.6612
3.9742 | -98.6471
3.9778 | -98.6348
3.9815 | -98.6336
3.9852 | -98.6307
3.9888 | -98.6435
3.9925 | -98.6639
3.9962 | -98.684
3.9998 | -98.6893
4.0035 | -98.6788
4.0071 | -98.6657
4.0108 | -98.6565
4.0145 | -98.633
4.0181 | -98.6023
4.0218 | -98.5665
4.0255 | -98.5294
4.0291 | -98.4936
4.0328 | -98.4572
4.0364 | -98.4159
4.0401 | -98.3704
4.0438 | -98.3455
4.0474 | -98.3472
4.0511 | -98.3682
4.0548 | -98.4133
4.0584 | -98.4884
4.0621 | -98.5796
4.0657 | -98.6716
4.0694 | -98.7566
4.0731 | -98.8252
4.0767 | -98.8864
4.0804 | -98.9338
4.0841 | -98.9578
4.0877 | -98.9645
4.0914 | -98.9576
4.0951 | -98.9297
4.0987 | -98.8862
4.1024 | -98.8376
4.106 | -98.786
4.1097 | -98.753
4.1134 | -98.7481
4.117 | -98.7646
4.1207 | -98.7904
4.1244 | -98.8243
4.128 | -98.8638
4.1317 | -98.899
4.1353 | -98.9377
4.139 | -98.9756
4.1427 | -99.0165
4.1463 | -99.0552
4.15 | -99.0929
4.1537 | -99.1172
4.1573 | -99.1164
4.161 | -99.0783
4.1646 | -98.9975
4.1683 | -98.8757
4.172 | -98.7201
4.1756 | -98.5549
4.1793 | -98.4041
4.183 | -98.2786
4.1866 | -98.1932
4.1903 | -98.1581
4.1939 | -98.181
4.1976 | -98.2458
4.2013 | -98.344
4.2049 | -98.4437
4.2086 | -98.5251
4.2123 | -98.5562
4.2159 | -98.5464
4.2196 | -98.5083
4.2233 | -98.4568
4.2269 | -98.4424
4.2306 | -98.429
4.2342 | -98.4378
4.2379 | -98.4632
4.2416 | -98.4969
4.2452 | -98.5389
4.2489 | -98.588
4.2526 | -98.6338
4.2562 | -98.6746
4.2599 | -98.712
4.2635 | -98.7449
4.2672 | -98.7779
4.2709 | -98.814
4.2745 | -98.8503
4.2782 | -98.8836
4.2819 | -98.9142
4.2855 | -98.94
4.2892 | -98.9603
4.2928 | -98.9771
4.2965 | -98.9856
4.3002 | -98.9851
4.3038 | -98.979
4.3075 | -98.9706
4.3112 | -98.9658
4.3148 | -98.9641
4.3185 | -98.9738
4.3221 | -99.0007
4.3258 | -99.034
4.3295 | -99.0763
4.3331 | -99.1335
4.3368 | -99.1961
4.3405 | -99.2645
4.3441 | -99.3231
4.3478 | -99.3772
4.3514 | -99.4148
4.3551 | -99.4345
4.3588 | -99.4326
4.3624 | -99.4036
4.3661 | -99.3513
4.3698 | -99.2734
4.3734 | -99.1755
4.3771 | -99.0767
4.3808 | -98.9791
4.3844 | -98.8952
4.3881 | -98.8328
4.3917 | -98.7997
4.3954 | -98.8084
4.3991 | -98.8529
4.4027 | -98.9205
4.4064 | -99.0056
4.4101 | -99.1036
4.4137 | -99.2002
4.4174 | -99.2767
4.421 | -99.3219
4.4247 | -99.3339
4.4284 | -99.3278
4.432 | -99.3085
4.4357 | -99.2847
4.4394 | -99.265
4.443 | -99.2533
4.4467 | -99.253
4.4503 | -99.2609
4.454 | -99.2727
4.4577 | -99.2905
4.4613 | -99.3062
4.465 | -99.3157
4.4687 | -99.3258
4.4723 | -99.3372
4.476 | -99.3509
4.4796 | -99.3619
4.4833 | -99.3779
4.487 | -99.3887
4.4906 | -99.3886
4.4943 | -99.3848
4.498 | -99.3751
4.5016 | -99.3599
4.5053 | -99.3449
4.509 | -99.3238
4.5126 | -99.2954
4.5163 | -99.2613
4.5199 | -99.2286
4.5236 | -99.1957
4.5273 | -99.1633
4.5309 | -99.1502
4.5346 | -99.1558
4.5383 | -99.18
4.5419 | -99.2287
4.5456 | -99.2918
4.5492 | -99.3652
4.5529 | -99.4336
4.5566 | -99.4967
4.5602 | -99.5636
4.5639 | -99.62
4.5676 | -99.671
4.5712 | -99.7242
4.5749 | -99.7873
4.5785 | -99.8588
4.5822 | -99.9269
4.5859 | -99.9823
4.5895 | -100.0186
4.5932 | -100.0191
4.5969 | -99.9798
4.6005 | -99.9056
4.6042 | -99.799
4.6078 | -99.6665
4.6115 | -99.526
4.6152 | -99.3857
4.6188 | -99.2615
4.6225 | -99.1575
4.6262 | -99.0872
4.6298 | -99.0431
4.6335 | -99.0264
4.6372 | -99.0403
4.6408 | -99.0822
4.6445 | -99.1401
4.6481 | -99.2061
4.6518 | -99.2822
4.6555 | -99.356
4.6591 | -99.4164
4.6628 | -99.47
4.6665 | -99.5207
4.6701 | -99.5652
4.6738 | -99.6169
4.6774 | -99.6827
4.6811 | -99.7504
4.6848 | -99.8212
4.6884 | -99.8812
4.6921 | -99.9224
4.6958 | -99.9443
4.6994 | -99.9494
4.7031 | -99.9415
4.7067 | -99.9056
4.7104 | -99.8536
4.7141 | -99.791
4.7177 | -99.7206
4.7214 | -99.6393
4.7251 | -99.5682
4.7287 | -99.5171
4.7324 | -99.4905
4.736 | -99.4858
4.7397 | -99.5058
4.7434 | -99.551
4.747 | -99.6166
4.7507 | -99.7022
4.7544 | -99.7978
4.758 | -99.8876
4.7617 | -99.9813
4.7653 | -100.0645
4.769 | -100.1186
4.7727 | -100.158
4.7763 | -100.1827
4.78 | -100.1913
4.7837 | -100.1879
4.7873 | -100.1841
4.791 | -100.1929
4.7947 | -100.2184
4.7983 | -100.2472
4.802 | -100.2778
4.8056 | -100.3145
4.8093 | -100.3476
4.813 | -100.3789
4.8166 | -100.3906
4.8203 | -100.3748
4.824 | -100.3489
4.8276 | -100.3155
4.8313 | -100.2599
4.8349 | -100.1926
4.8386 | -100.1194
4.8423 | -100.0456
4.8459 | -99.971
4.8496 | -99.8945
4.8533 | -99.8286
4.8569 | -99.7671
4.8606 | -99.711
4.8642 | -99.6711
4.8679 | -99.6441
4.8716 | -99.6222
4.8752 | -99.6065
4.8789 | -99.5967
4.8826 | -99.5816
4.8862 | -99.5689
4.8899 | -99.5566
4.8935 | -99.5647
4.8972 | -99.5932
4.9009 | -99.632
4.9045 | -99.6843
4.9082 | -99.7478
4.9119 | -99.8214
4.9155 | -99.8975
4.9192 | -99.973
4.9229 | -100.0401
4.9265 | -100.1073
4.9302 | -100.1671
4.9338 | -100.2292
4.9375 | -100.2826
4.9412 | -100.3167
4.9448 | -100.3387
4.9485 | -100.3493
4.9522 | -100.3517
4.9558 | -100.3401
4.9595 | -100.3235
4.9631 | -100.305
4.9668 | -100.2906
4.9705 | -100.2759
4.9741 | -100.2658
4.9778 | -100.2612
4.9815 | -100.2616
4.9851 | -100.2699
4.9888 | -100.2664
4.9924 | -100.2614
4.9961 | -100.2548
4.9998 | -100.2473
5.0034 | -100.2271
5.0071 | -100.2049
5.0108 | -100.1831
5.0144 | -100.1592
5.0181 | -100.1412
5.0217 | -100.1227
5.0254 | -100.1013
5.0291 | -100.0742
5.0327 | -100.0473
5.0364 | -100.0166
5.0401 | -99.9771
5.0437 | -99.929
5.0474 | -99.8801
5.0511 | -99.8332
5.0547 | -99.7929
5.0584 | -99.7644
5.062 | -99.7462
5.0657 | -99.7384
5.0694 | -99.7346
5.073 | -99.7393
5.0767 | -99.7452
5.0804 | -99.7565
5.084 | -99.7712
5.0877 | -99.7883
5.0913 | -99.8144
5.095 | -99.8429
5.0987 | -99.8817
5.1023 | -99.9252
5.106 | -99.9701
5.1097 | -100.0138
5.1133 | -100.0516
5.117 | -100.0963
5.1206 | -100.1374
5.1243 | -100.1811
5.128 | -100.2209
5.1316 | -100.2501
5.1353 | -100.2644
5.139 | -100.2585
5.1426 | -100.2421
5.1463 | -100.2073
5.1499 | -100.1544
5.1536 | -100.0996
5.1573 | -100.0501
5.1609 | -100.0106
5.1646 | -99.9856
5.1683 | -99.9715
5.1719 | -99.9722
5.1756 | -99.9738
5.1792 | -99.9708
5.1829 | -99.9686
5.1866 | -99.9627
5.1902 | -99.9631
5.1939 | -99.9629
5.1976 | -99.9738
5.2012 | -99.9791
5.2049 | -99.9896
5.2086 | -100.0105
5.2122 | -100.032
5.2159 | -100.048
5.2195 | -100.0532
5.2232 | -100.0606
5.2269 | -100.0722
5.2305 | -100.0864
5.2342 | -100.1033
5.2379 | -100.1339
5.2415 | -100.1705
5.2452 | -100.2188
5.2488 | -100.2719
5.2525 | -100.3346
5.2562 | -100.4016
5.2598 | -100.4538
5.2635 | -100.4929
5.2672 | -100.5183
5.2708 | -100.5388
5.2745 | -100.5496
5.2781 | -100.5602
5.2818 | -100.5507
5.2855 | -100.5453
5.2891 | -100.5194
5.2928 | -100.4894
5.2965 | -100.4567
5.3001 | -100.4111
5.3038 | -100.3556
5.3074 | -100.3054
5.3111 | -100.2639
5.3148 | -100.2257
5.3184 | -100.191
5.3221 | -100.1654
5.3258 | -100.1378
5.3294 | -100.1085
5.3331 | -100.0813
5.3368 | -100.0521
5.3404 | -100.0297
5.3441 | -100.0194
5.3477 | -100.0207
5.3514 | -100.0316
5.3551 | -100.0551
5.3587 | -100.0926
5.3624 | -100.138
5.3661 | -100.192
5.3697 | -100.2559
5.3734 | -100.3259
5.377 | -100.3915
5.3807 | -100.4538
5.3844 | -100.5089
5.388 | -100.5459
5.3917 | -100.5627
5.3954 | -100.5681
5.399 | -100.5694
5.4027 | -100.567
5.4063 | -100.5676
5.41 | -100.5742
5.4137 | -100.5878
5.4173 | -100.6027
5.421 | -100.6241
5.4247 | -100.6381
5.4283 | -100.6482
5.432 | -100.6566
5.4356 | -100.6659
5.4393 | -100.6771
5.443 | -100.6944
5.4466 | -100.7163
5.4503 | -100.7451
5.454 | -100.7781
5.4576 | -100.8064
5.4613 | -100.8317
5.465 | -100.853
5.4686 | -100.8692
5.4723 | -100.8754
5.4759 | -100.8668
5.4796 | -100.8414
5.4833 | -100.8059
5.4869 | -100.7609
5.4906 | -100.7097
5.4943 | -100.6579
5.4979 | -100.6019
5.5016 | -100.5491
5.5052 | -100.5064
5.5089 | -100.4726
5.5126 | -100.4451
5.5162 | -100.4297
5.5199 | -100.4213
5.5236 | -100.4232
5.5272 | -100.4363
5.5309 | -100.4608
5.5345 | -100.4944
5.5382 | -100.5301
5.5419 | -100.5741
5.5455 | -100.6163
5.5492 | -100.6553
5.5529 | -100.6959
5.5565 | -100.7359
5.5602 | -100.7642
5.5638 | -100.7849
5.5675 | -100.7886
5.5712 | -100.7815
5.5748 | -100.7695
5.5785 | -100.7561
5.5822 | -100.7465
5.5858 | -100.7303
5.5895 | -100.7104
5.5931 | -100.6922
5.5968 | -100.6656
5.6005 | -100.6237
5.6041 | -100.5692
5.6078 | -100.5001
5.6115 | -100.4116
5.6151 | -100.3239
5.6188 | -100.237
5.6225 | -100.1503
5.6261 | -100.0695
5.6298 | -100.0099
5.6334 | -99.9653
5.6371 | -99.9379
5.6408 | -99.938
5.6444 | -99.9622
5.6481 | -100.01
5.6518 | -100.0711
5.6554 | -100.1381
5.6591 | -100.2033
5.6627 | -100.2647
5.6664 | -100.3192
5.6701 | -100.3623
5.6737 | -100.3869
5.6774 | -100.4045
5.6811 | -100.4233
5.6847 | -100.4449
5.6884 | -100.4729
5.692 | -100.5044
5.6957 | -100.5345
5.6994 | -100.5512
5.703 | -100.5633
5.7067 | -100.5715
5.7104 | -100.568
5.714 | -100.5472
5.7177 | -100.5326
5.7214 | -100.5203
5.725 | -100.5033
5.7287 | -100.5034
5.7323 | -100.5024
5.736 | -100.5069
5.7397 | -100.5037
5.7433 | -100.5008
5.747 | -100.4922
5.7507 | -100.4702
5.7543 | -100.4425
5.758 | -100.4078
5.7616 | -100.3575
5.7653 | -100.3004
5.769 | -100.2444
5.7726 | -100.1889
5.7763 | -100.1448
5.78 | -100.1106
5.7836 | -100.0828
5.7873 | -100.0763
5.7909 | -100.079
5.7946 | -100.0868
5.7983 | -100.1002
5.8019 | -100.1111
5.8056 | -100.1233
5.8093 | -100.1296
5.8129 | -100.1404
5.8166 | -100.1442
5.8202 | -100.1534
5.8239 | -100.1667
5.8276 | -100.1818
5.8312 | -100.189
5.8349 | -100.1886
5.8386 | -100.1927
5.8422 | -100.1908
5.8459 | -100.1786
5.8495 | -100.1585
5.8532 | -100.1271
5.8569 | -100.0892
5.8605 | -100.0525
5.8642 | -100.014
5.8679 | -99.9825
5.8715 | -99.9506
5.8752 | -99.9108
5.8789 | -99.8783
5.8825 | -99.8511
5.8862 | -99.8241
5.8898 | -99.7967
5.8935 | -99.7731
5.8972 | -99.7573
5.9008 | -99.734
5.9045 | -99.7122
5.9082 | -99.7007
5.9118 | -99.694
5.9155 | -99.6916
5.9191 | -99.6895
5.9228 | -99.6916
5.9265 | -99.707
5.9301 | -99.7225
5.9338 | -99.7389
5.9375 | -99.7595
5.9411 | -99.7799
5.9448 | -99.8039
5.9484 | -99.8324
5.9521 | -99.8687
5.9558 | -99.9072
5.9594 | -99.9518
5.9631 | -100.0093
5.9668 | -100.0665
5.9704 | -100.1282
5.9741 | -100.1934
5.9777 | -100.2608
5.9814 | -100.3216
5.9851 | -100.3785
5.9887 | -100.4366
5.9924 | -100.4952
5.9961 | -100.5471
5.9997 | -100.5825
6.0034 | -100.6115
6.0071 | -100.6374
6.0107 | -100.6659
6.0144 | -100.6865
6.018 | -100.6886
6.0217 | -100.6958
6.0254 | -100.7076
6.029 | -100.7165
6.0327 | -100.7245
6.0364 | -100.7329
6.04 | -100.7415
6.0437 | -100.746
6.0473 | -100.763
6.051 | -100.7801
6.0547 | -100.7945
6.0583 | -100.8123
6.062 | -100.8303
6.0657 | -100.851
6.0693 | -100.872
6.073 | -100.8986
6.0766 | -100.9241
6.0803 | -100.9398
6.084 | -100.9587
6.0876 | -100.9726
6.0913 | -100.9824
6.095 | -100.9818
6.0986 | -100.9768
6.1023 | -100.9688
6.1059 | -100.9585
6.1096 | -100.9573
6.1133 | -100.9478
6.1169 | -100.944
6.1206 | -100.9482
6.1243 | -100.9634
6.1279 | -100.9799
6.1316 | -100.9985
6.1353 | -101.0223
6.1389 | -101.0377
6.1426 | -101.0558
6.1462 | -101.0729
6.1499 | -101.0845
6.1536 | -101.0937
6.1572 | -101.0982
6.1609 | -101.1049
6.1646 | -101.1184
6.1682 | -101.1308
6.1719 | -101.1476
6.1755 | -101.1704
6.1792 | -101.1975
6.1829 | -101.2254
6.1865 | -101.2613
6.1902 | -101.2902
6.1939 | -101.3179
6.1975 | -101.3501
6.2012 | -101.3801
6.2048 | -101.4112
6.2085 | -101.4386
6.2122 | -101.4651
6.2158 | -101.4815
6.2195 | -101.4965
6.2232 | -101.5061
6.2268 | -101.502
6.2305 | -101.4955
6.2341 | -101.4809
6.2378 | -101.4627
6.2415 | -101.4433
6.2451 | -101.4242
6.2488 | -101.4064
6.2525 | -101.3878
6.2561 | -101.3655
6.2598 | -101.3463
6.2634 | -101.3332
6.2671 | -101.3218
6.2708 | -101.3191
6.2744 | -101.3245
6.2781 | -101.3321
6.2818 | -101.3549
6.2854 | -101.3851
6.2891 | -101.4164
6.2928 | -101.4514
6.2964 | -101.4891
6.3001 | -101.5234
6.3037 | -101.558
6.3074 | -101.585
6.3111 | -101.5978
6.3147 | -101.6109
6.3184 | -101.6273
6.3221 | -101.6359
6.3257 | -101.6416
6.3294 | -101.6485
6.333 | -101.6595
6.3367 | -101.6652
6.3404 | -101.6608
6.344 | -101.664
6.3477 | -101.6574
6.3514 | -101.6495
6.355 | -101.6405
6.3587 | -101.6257
6.3623 | -101.6075
6.366 | -101.5892
6.3697 | -101.5783
6.3733 | -101.5712
6.377 | -101.5605
6.3807 | -101.5515
6.3843 | -101.5444
6.388 | -101.5384
6.3916 | -101.5386
6.3953 | -101.5406
6.399 | -101.5463
6.4026 | -101.5614
6.4063 | -101.5838
6.41 | -101.6092
6.4136 | -101.6391
6.4173 | -101.6782
6.421 | -101.719
6.4246 | -101.7545
6.4283 | -101.7838
6.4319 | -101.8137
6.4356 | -101.8374
6.4393 | -101.8549
6.4429 | -101.8606
6.4466 | -101.8652
6.4503 | -101.8565
6.4539 | -101.8387
6.4576 | -101.824
6.4612 | -101.8102
6.4649 | -101.794
6.4686 | -101.7684
6.4722 | -101.7451
6.4759 | -101.7268
6.4796 | -101.7039
6.4832 | -101.6815
6.4869 | -101.659
6.4905 | -101.6321
6.4942 | -101.6128
6.4979 | -101.5953
6.5015 | -101.5864
6.5052 | -101.5822
6.5089 | -101.5832
6.5125 | -101.5928
6.5162 | -101.6085
6.5198 | -101.6308
6.5235 | -101.6585
6.5272 | -101.6804
6.5308 | -101.694
6.5345 | -101.7061
6.5382 | -101.7128
6.5418 | -101.707
6.5455 | -101.6976
6.5492 | -101.6944
6.5528 | -101.6952
6.5565 | -101.6894
6.5601 | -101.6908
6.5638 | -101.7074
6.5675 | -101.734
6.5711 | -101.7665
6.5748 | -101.8074
6.5785 | -101.8583
6.5821 | -101.915
6.5858 | -101.9736
6.5894 | -102.0346
6.5931 | -102.0859
6.5968 | -102.1314
6.6004 | -102.1661
6.6041 | -102.1958
6.6078 | -102.2109
6.6114 | -102.2108
6.6151 | -102.1975
6.6187 | -102.1738
6.6224 | -102.1394
6.6261 | -102.0948
6.6297 | -102.0489
6.6334 | -101.9938
6.6371 | -101.9446
6.6407 | -101.8993
6.6444 | -101.8546
6.648 | -101.8075
6.6517 | -101.7638
6.6554 | -101.7354
6.659 | -101.7159
6.6627 | -101.702
6.6664 | -101.6982
6.67 | -101.7057
6.6737 | -101.7221
6.6773 | -101.7465
6.681 | -101.7761
6.6847 | -101.8105
6.6883 | -101.8481
6.692 | -101.8864
6.6957 | -101.9208
6.6993 | -101.9526
6.703 | -101.9722
6.7067 | -101.9855
6.7103 | -101.9933
6.714 | -102.0003
6.7176 | -102.0061
6.7213 | -102.0118
6.725 | -102.0218
6.7286 | -102.0276
6.7323 | -102.0438
6.736 | -102.0615
6.7396 | -102.0931
6.7433 | -102.1343
6.7469 | -102.1834
6.7506 | -102.2355
6.7543 | -102.2774
6.7579 | -102.3227
6.7616 | -102.3601
6.7653 | -102.3973
6.7689 | -102.4284
6.7726 | -102.4459
6.7762 | -102.4454
6.7799 | -102.4272
6.7836 | -102.3962
6.7872 | -102.352
6.7909 | -102.3038
6.7946 | -102.2585
6.7982 | -102.2139
6.8019 | -102.1854
6.8055 | -102.1688
6.8092 | -102.166
6.8129 | -102.17
6.8165 | -102.1872
6.8202 | -102.2121
6.8239 | -102.2344
6.8275 | -102.2531
6.8312 | -102.2622
6.8349 | -102.2665
6.8385 | -102.2613
6.8422 | -102.2477
6.8458 | -102.2307
6.8495 | -102.2102
6.8532 | -102.1948
6.8568 | -102.1736
6.8605 | -102.16
6.8642 | -102.1436
6.8678 | -102.1523
6.8715 | -102.1416
6.8751 | -102.1385
6.8788 | -102.1465
6.8825 | -102.1596
6.8861 | -102.1758
6.8898 | -102.1955
6.8935 | -102.2107
6.8971 | -102.2255
6.9008 | -102.2461
6.9044 | -102.261
6.9081 | -102.2712
6.9118 | -102.2776
6.9154 | -102.2827
6.9191 | -102.2825
6.9228 | -102.2768
6.9264 | -102.2711
6.9301 | -102.265
6.9337 | -102.2576
6.9374 | -102.2464
6.9411 | -102.2362
6.9447 | -102.2229
6.9484 | -102.2141
6.9521 | -102.2041
6.9557 | -102.1852
6.9594 | -102.1649
6.9631 | -102.1461
6.9667 | -102.1287
6.9704 | -102.1157
6.974 | -102.1095
6.9777 | -102.1112
6.9814 | -102.1179
6.985 | -102.1257
6.9887 | -102.1386
6.9924 | -102.1536
6.996 | -102.1699
6.9997 | -102.1902
7.0033 | -102.2021
7.007 | -102.2062
7.0107 | -102.2072
7.0143 | -102.2053
7.018 | -102.1927
7.0217 | -102.1743
7.0253 | -102.1535
7.029 | -102.1322
7.0326 | -102.1083
7.0363 | -102.076
7.04 | -102.0514
7.0436 | -102.0301
7.0473 | -102.0115
7.051 | -101.9992
7.0546 | -101.9911
7.0583 | -101.9949
7.0619 | -102.0077
7.0656 | -102.0227
7.0693 | -102.0385
7.0729 | -102.061
7.0766 | -102.0884
7.0803 | -102.1119
7.0839 | -102.1414
7.0876 | -102.1769
7.0912 | -102.2195
7.0949 | -102.2654
7.0986 | -102.312
7.1022 | -102.3593
7.1059 | -102.4108
7.1096 | -102.4644
7.1132 | -102.5171
7.1169 | -102.569
7.1206 | -102.6168
7.1242 | -102.6695
7.1279 | -102.7173
7.1315 | -102.7613
7.1352 | -102.7994
7.1389 | -102.8346
7.1425 | -102.8663
7.1462 | -102.8757
7.1499 | -102.8757
7.1535 | -102.8611
7.1572 | -102.8343
7.1608 | -102.7951
7.1645 | -102.7475
7.1682 | -102.6984
7.1718 | -102.6485
7.1755 | -102.6001
7.1792 | -102.5529
7.1828 | -102.5115
7.1865 | -102.4805
7.1901 | -102.4572
7.1938 | -102.4472
7.1975 | -102.4422
7.2011 | -102.4487
7.2048 | -102.4735
7.2085 | -102.5079
7.2121 | -102.541
7.2158 | -102.5809
7.2195 | -102.6219
7.2231 | -102.6622
7.2268 | -102.7013
7.2304 | -102.7267
7.2341 | -102.7458
7.2378 | -102.7559
7.2414 | -102.7591
7.2451 | -102.7547
7.2488 | -102.7428
7.2524 | -102.7238
7.2561 | -102.6989
7.2597 | -102.6715
7.2634 | -102.6429
7.2671 | -102.6097
7.2707 | -102.5762
7.2744 | -102.5405
7.2781 | -102.5083
7.2817 | -102.4821
7.2854 | -102.4584
7.289 | -102.4387
7.2927 | -102.4236
7.2964 | -102.4126
7.3 | -102.4044
7.3037 | -102.4077
7.3074 | -102.4155
7.311 | -102.4244
7.3147 | -102.4463
7.3183 | -102.4823
7.322 | -102.5275
7.3257 | -102.5788
7.3293 | -102.6332
7.333 | -102.689
7.3367 | -102.7537
7.3403 | -102.8153
7.344 | -102.8758
7.3476 | -102.9358
7.3513 | -102.9852
7.355 | -103.0264
7.3586 | -103.0648
7.3623 | -103.0948
7.366 | -103.1162
7.3696 | -103.1239
7.3733 | -103.1243
7.377 | -103.1188
7.3806 | -103.1082
7.3843 | -103.0851
7.3879 | -103.0529
7.3916 | -103.0073
7.3953 | -102.9411
7.3989 | -102.8903
7.4026 | -102.8428
7.4063 | -102.7935
7.4099 | -102.7441
7.4136 | -102.7011
7.4172 | -102.6644
7.4209 | -102.634
7.4246 | -102.607
7.4282 | -102.5823
7.4319 | -102.5672
7.4356 | -102.5632
7.4392 | -102.5672
7.4429 | -102.5837
7.4465 | -102.6062
7.4502 | -102.6365
7.4539 | -102.673
7.4575 | -102.7129
7.4612 | -102.7608
7.4649 | -102.8149
7.4685 | -102.8661
7.4722 | -102.9125
7.4758 | -102.9576
7.4795 | -103.0035
7.4832 | -103.0447
7.4868 | -103.0853
7.4905 | -103.1266
7.4942 | -103.1672
7.4978 | -103.2074
7.5015 | -103.2493
7.5052 | -103.2862
7.5088 | -103.312
7.5125 | -103.3312
7.5161 | -103.3465
7.5198 | -103.3579
7.5235 | -103.366
7.5271 | -103.3712
7.5308 | -103.3668
7.5345 | -103.3597
7.5381 | -103.3562
7.5418 | -103.3529
7.5454 | -103.3455
7.5491 | -103.3348
7.5528 | -103.3211
7.5564 | -103.3008
7.5601 | -103.2798
7.5638 | -103.2618
7.5674 | -103.2418
7.5711 | -103.2171
7.5747 | -103.1949
7.5784 | -103.1728
7.5821 | -103.1523
7.5857 | -103.1366
7.5894 | -103.1227
7.5931 | -103.1127
7.5967 | -103.1064
7.6004 | -103.0951
7.604 | -103.0849
7.6077 | -103.0792
7.6114 | -103.0713
7.615 | -103.0682
7.6187 | -103.0701
7.6224 | -103.0718
7.626 | -103.0679
7.6297 | -103.0611
7.6334 | -103.0534
7.637 | -103.0444
7.6407 | -103.0335
7.6443 | -103.0239
7.648 | -103.0116
7.6517 | -102.9957
7.6553 | -102.9776
7.659 | -102.9591
7.6627 | -102.9425
7.6663 | -102.9307
7.67 | -102.9247
7.6736 | -102.9176
7.6773 | -102.9103
7.681 | -102.9077
7.6846 | -102.9087
7.6883 | -102.9001
7.692 | -102.8895
7.6956 | -102.8782
7.6993 | -102.8655
7.7029 | -102.8548
7.7066 | -102.8428
7.7103 | -102.8325
7.7139 | -102.8244
7.7176 | -102.815
7.7213 | -102.8039
7.7249 | -102.792
7.7286 | -102.7789
7.7322 | -102.7644
7.7359 | -102.7531
7.7396 | -102.7537
7.7432 | -102.7554
7.7469 | -102.7625
7.7506 | -102.7773
7.7542 | -102.7969
7.7579 | -102.8231
7.7615 | -102.8541
7.7652 | -102.8888
7.7689 | -102.9291
7.7725 | -102.9674
7.7762 | -103.0004
7.7799 | -103.0234
7.7835 | -103.0395
7.7872 | -103.0558
7.7909 | -103.0664
7.7945 | -103.0718
7.7982 | -103.0788
7.8018 | -103.0867
7.8055 | -103.0933
7.8092 | -103.0995
7.8128 | -103.1054
7.8165 | -103.1083
7.8202 | -103.116
7.8238 | -103.1232
7.8275 | -103.1357
7.8311 | -103.1522
7.8348 | -103.1745
7.8385 | -103.1988
7.8421 | -103.2221
7.8458 | -103.2478
7.8495 | -103.2747
7.8531 | -103.2908
7.8568 | -103.3053
7.8604 | -103.3206
7.8641 | -103.3278
7.8678 | -103.3318
7.8714 | -103.336
7.8751 | -103.3338
7.8788 | -103.3268
7.8824 | -103.32
7.8861 | -103.314
7.8897 | -103.3118
7.8934 | -103.3139
7.8971 | -103.3156
7.9007 | -103.3135
7.9044 | -103.3151
7.9081 | -103.3197
7.9117 | -103.3252
7.9154 | -103.3334
7.9191 | -103.3421
7.9227 | -103.3539
7.9264 | -103.3625
7.93 | -103.367
7.9337 | -103.3655
7.9374 | -103.3633
7.941 | -103.3617
7.9447 | -103.3567
7.9484 | -103.3484
7.952 | -103.3275
7.9557 | -103.3023
7.9593 | -103.2755
7.963 | -103.2483
7.9667 | -103.2202
7.9703 | -103.1989
7.974 | -103.1806
7.9777 | -103.1615
7.9813 | -103.1491
7.985 | -103.1425
7.9886 | -103.1385
7.9923 | -103.1432
7.996 | -103.151
7.9996 | -103.1621
8.0033 | -103.1775
8.007 | -103.1904
8.0106 | -103.2005
8.0143 | -103.208
8.0179 | -103.2167
8.0216 | -103.2151
8.0253 | -103.2074
8.0289 | -103.1955
8.0326 | -103.1839
8.0363 | -103.1719
8.0399 | -103.1581
8.0436 | -103.1432
8.0473 | -103.1323
8.0509 | -103.1248
8.0546 | -103.1243
8.0582 | -103.1264
8.0619 | -103.1297
8.0656 | -103.134
8.0692 | -103.1474
8.0729 | -103.1731
8.0766 | -103.2064
8.0802 | -103.2427
8.0839 | -103.2787
8.0875 | -103.3176
8.0912 | -103.3543
8.0949 | -103.3834
8.0985 | -103.4067
8.1022 | -103.4243
8.1059 | -103.4368
8.1095 | -103.4418
8.1132 | -103.4348
8.1168 | -103.4217
8.1205 | -103.4021
8.1242 | -103.3807
8.1278 | -103.3601
8.1315 | -103.3433
8.1352 | -103.3259
8.1388 | -103.3029
8.1425 | -103.2877
8.1461 | -103.2768
8.1498 | -103.2658
8.1535 | -103.2582
8.1571 | -103.2554
8.1608 | -103.2602
8.1645 | -103.2723
8.1681 | -103.2868
8.1718 | -103.3063
8.1754 | -103.3276
8.1791 | -103.3494
8.1828 | -103.3734
8.1864 | -103.4014
8.1901 | -103.4184
8.1938 | -103.4331
8.1974 | -103.4393
8.2011 | -103.4373
8.2048 | -103.4368
8.2084 | -103.4323
8.2121 | -103.4214
8.2157 | -103.4141
8.2194 | -103.4021
8.2231 | -103.3791
8.2267 | -103.3491
8.2304 | -103.3254
8.2341 | -103.3073
8.2377 | -103.2872
8.2414 | -103.2685
8.245 | -103.2528
8.2487 | -103.244
8.2524 | -103.2391
8.256 | -103.2429
8.2597 | -103.2582
8.2634 | -103.273
8.267 | -103.284
8.2707 | -103.3003
8.2743 | -103.3158
8.278 | -103.3316
8.2817 | -103.3459
8.2853 | -103.3566
8.289 | -103.3653
8.2927 | -103.3696
8.2963 | -103.3725
8.3 | -103.3738
8.3036 | -103.3718
8.3073 | -103.3636
8.311 | -103.3594
8.3146 | -103.3448
8.3183 | -103.3264
8.322 | -103.3074
8.3256 | -103.2826
8.3293 | -103.2601
8.333 | -103.2386
8.3366 | -103.2176
8.3403 | -103.2035
8.3439 | -103.1942
8.3476 | -103.1876
8.3513 | -103.1799
8.3549 | -103.1719
8.3586 | -103.1713
8.3623 | -103.1726
8.3659 | -103.1748
8.3696 | -103.1783
8.3732 | -103.1797
8.3769 | -103.1835
8.3806 | -103.1845
8.3842 | -103.1849
8.3879 | -103.1849
8.3916 | -103.1815
8.3952 | -103.1737
8.3989 | -103.1641
8.4025 | -103.1493
8.4062 | -103.1325
8.4099 | -103.1234
8.4135 | -103.1167
8.4172 | -103.1152
8.4209 | -103.1155
8.4245 | -103.1243
8.4282 | -103.1381
8.4318 | -103.1496
8.4355 | -103.1672
8.4392 | -103.1902
8.4428 | -103.209
8.4465 | -103.2201
8.4502 | -103.2257
8.4538 | -103.2259
8.4575 | -103.232
8.4612 | -103.2357
8.4648 | -103.2337
8.4685 | -103.2381
8.4721 | -103.2452
8.4758 | -103.2482
8.4795 | -103.2548
8.4831 | -103.2677
8.4868 | -103.2847
8.4905 | -103.3005
8.4941 | -103.3187
8.4978 | -103.3421
8.5014 | -103.365
8.5051 | -103.3896
8.5088 | -103.4153
8.5124 | -103.4424
8.5161 | -103.4665
8.5198 | -103.4902
8.5234 | -103.5149
8.5271 | -103.5421
8.5307 | -103.5693
8.5344 | -103.589
8.5381 | -103.6101
8.5417 | -103.6318
8.5454 | -103.6541
8.5491 | -103.6733
8.5527 | -103.6905
8.5564 | -103.7098
8.56 | -103.7251
8.5637 | -103.732
8.5674 | -103.7414
8.571 | -103.751
8.5747 | -103.7574
8.5784 | -103.7574
8.582 | -103.7537
8.5857 | -103.7493
8.5893 | -103.7493
8.593 | -103.7548
8.5967 | -103.7601
8.6003 | -103.7716
8.604 | -103.7897
8.6077 | -103.8039
8.6113 | -103.8155
8.615 | -103.8289
8.6187 | -103.8435
8.6223 | -103.8526
8.626 | -103.8607
8.6296 | -103.8684
8.6333 | -103.8691
8.637 | -103.8674
8.6406 | -103.863
8.6443 | -103.8544
8.648 | -103.8426
8.6516 | -103.829
8.6553 | -103.8178
8.6589 | -103.8068
8.6626 | -103.7964
8.6663 | -103.7895
8.6699 | -103.7871
8.6736 | -103.7914
8.6773 | -103.7995
8.6809 | -103.8127
8.6846 | -103.8293
8.6882 | -103.8486
8.6919 | -103.8739
8.6956 | -103.8964
8.6992 | -103.9143
8.7029 | -103.9349
8.7066 | -103.9563
8.7102 | -103.9679
8.7139 | -103.9738
8.7175 | -103.9699
8.7212 | -103.957
8.7249 | -103.9342
8.7285 | -103.8992
8.7322 | -103.8594
8.7359 | -103.8215
8.7395 | -103.7736
8.7432 | -103.721
8.7469 | -103.6696
8.7505 | -103.6204
8.7542 | -103.5707
8.7578 | -103.5296
8.7615 | -103.4983
8.7652 | -103.4704
8.7688 | -103.4509
8.7725 | -103.4412
8.7762 | -103.4399
8.7798 | -103.4413
8.7835 | -103.4464
8.7871 | -103.4561
8.7908 | -103.47
8.7945 | -103.4809
8.7981 | -103.4918
8.8018 | -103.5038
8.8055 | -103.5142
8.8091 | -103.5198
8.8128 | -103.5232
8.8164 | -103.5258
8.8201 | -103.5259
8.8238 | -103.5216
8.8274 | -103.5219
8.8311 | -103.5154
8.8348 | -103.5069
8.8384 | -103.5023
8.8421 | -103.5019
8.8457 | -103.4984
8.8494 | -103.4953
8.8531 | -103.4961
8.8567 | -103.4971
8.8604 | -103.5
8.8641 | -103.5029
8.8677 | -103.5122
8.8714 | -103.5238
8.8751 | -103.5354
8.8787 | -103.5505
8.8824 | -103.5693
8.886 | -103.5793
8.8897 | -103.5938
8.8934 | -103.6161
8.897 | -103.6388
8.9007 | -103.6628
8.9044 | -103.6883
8.908 | -103.7101
8.9117 | -103.7288
8.9153 | -103.7465
8.919 | -103.7599
8.9227 | -103.7697
8.9263 | -103.78
8.93 | -103.7887
8.9337 | -103.7958
8.9373 | -103.8018
8.941 | -103.8046
8.9446 | -103.8034
8.9483 | -103.803
8.952 | -103.8106
8.9556 | -103.812
8.9593 | -103.812
8.963 | -103.8148
8.9666 | -103.8149
8.9703 | -103.81
8.9739 | -103.808
8.9776 | -103.8116
8.9813 | -103.8201
8.9849 | -103.822
8.9886 | -103.8298
8.9923 | -103.8388
8.9959 | -103.8454
8.9996 | -103.8549
9.0033 | -103.8679
9.0069 | -103.8869
9.0106 | -103.8996
9.0142 | -103.9109
9.0179 | -103.9307
9.0216 | -103.9443
9.0252 | -103.9532
9.0289 | -103.9623
9.0326 | -103.9696
9.0362 | -103.9734
9.0399 | -103.9816
9.0435 | -103.9846
9.0472 | -103.9844
9.0509 | -103.9824
9.0545 | -103.981
9.0582 | -103.9713
9.0619 | -103.962
9.0655 | -103.9557
9.0692 | -103.9473
9.0728 | -103.941
9.0765 | -103.9346
9.0802 | -103.9288
9.0838 | -103.9232
9.0875 | -103.9219
9.0912 | -103.9233
9.0948 | -103.9234
9.0985 | -103.9305
9.1021 | -103.937
9.1058 | -103.9486
9.1095 | -103.9599
9.1131 | -103.9728
9.1168 | -103.99
9.1205 | -104.008
9.1241 | -104.0261
9.1278 | -104.0463
9.1314 | -104.0727
9.1351 | -104.1001
9.1388 | -104.1285
9.1424 | -104.1544
9.1461 | -104.186
9.1498 | -104.218
9.1534 | -104.2466
9.1571 | -104.2755
9.1608 | -104.3047
9.1644 | -104.3258
9.1681 | -104.3464
9.1717 | -104.3665
9.1754 | -104.3779
9.1791 | -104.3858
9.1827 | -104.3885
9.1864 | -104.3865
9.1901 | -104.3775
9.1937 | -104.3698
9.1974 | -104.3541
9.201 | -104.3369
9.2047 | -104.3189
9.2084 | -104.298
9.212 | -104.279
9.2157 | -104.2563
9.2194 | -104.2339
9.223 | -104.2184
9.2267 | -104.2067
9.2303 | -104.1964
9.234 | -104.1815
9.2377 | -104.1681
9.2413 | -104.1539
9.245 | -104.1428
9.2487 | -104.1353
9.2523 | -104.1232
9.256 | -104.106
9.2596 | -104.0868
9.2633 | -104.0678
9.267 | -104.0449
9.2706 | -104.0185
9.2743 | -103.9913
9.278 | -103.9717
9.2816 | -103.9503
9.2853 | -103.9254
9.289 | -103.8998
9.2926 | -103.8796
9.2963 | -103.8603
9.2999 | -103.8464
9.3036 | -103.8363
9.3073 | -103.8223
9.3109 | -103.8163
9.3146 | -103.8162
9.3183 | -103.8178
9.3219 | -103.8252
9.3256 | -103.8357
9.3292 | -103.842
9.3329 | -103.8562
9.3366 | -103.8727
9.3402 | -103.8964
9.3439 | -103.9258
9.3476 | -103.9557
9.3512 | -103.9909
9.3549 | -104.0229
9.3585 | -104.0527
9.3622 | -104.0851
9.3659 | -104.1208
9.3695 | -104.1554
9.3732 | -104.1901
9.3769 | -104.221
9.3805 | -104.2462
9.3842 | -104.2676
9.3878 | -104.2871
9.3915 | -104.3051
9.3952 | -104.3163
9.3988 | -104.3213
9.4025 | -104.3222
9.4062 | -104.3171
9.4098 | -104.3078
9.4135 | -104.292
9.4172 | -104.274
9.4208 | -104.2571
9.4245 | -104.2372
9.4281 | -104.2156
9.4318 | -104.1922
9.4355 | -104.1721
9.4391 | -104.1556
9.4428 | -104.1509
9.4465 | -104.1452
9.4501 | -104.1388
9.4538 | -104.1352
9.4574 | -104.1345
9.4611 | -104.1387
9.4648 | -104.1462
9.4684 | -104.1504
9.4721 | -104.1528
9.4758 | -104.1517
9.4794 | -104.1474
9.4831 | -104.141
9.4867 | -104.1285
9.4904 | -104.1213
9.4941 | -104.1145
9.4977 | -104.1044
9.5014 | -104.0895
9.5051 | -104.0805
9.5087 | -104.0663
9.5124 | -104.0559
9.516 | -104.0468
9.5197 | -104.0389
9.5234 | -104.0318
9.527 | -104.0324
9.5307 | -104.0358
9.5344 | -104.0426
9.538 | -104.0528
9.5417 | -104.0682
9.5453 | -104.0928
9.549 | -104.1226
9.5527 | -104.1551
9.5563 | -104.1983
9.56 | -104.2414
9.5637 | -104.2869
9.5673 | -104.3357
9.571 | -104.3848
9.5747 | -104.4305
9.5783 | -104.472
9.582 | -104.5124
9.5856 | -104.5518
9.5893 | -104.589
9.593 | -104.6191
9.5966 | -104.6408
9.6003 | -104.6558
9.604 | -104.6649
9.6076 | -104.6662
9.6113 | -104.6599
9.6149 | -104.6472
9.6186 | -104.6305
9.6223 | -104.6134
9.6259 | -104.5909
9.6296 | -104.5585
9.6333 | -104.5218
9.6369 | -104.4864
9.6406 | -104.4559
9.6442 | -104.4252
9.6479 | -104.3957
9.6516 | -104.3686
9.6552 | -104.3444
9.6589 | -104.3221
9.6626 | -104.3006
9.6662 | -104.2804
9.6699 | -104.2581
9.6736 | -104.2404
9.6772 | -104.2237
9.6809 | -104.2111
9.6845 | -104.2036
9.6882 | -104.1965
9.6919 | -104.1919
9.6955 | -104.1917
9.6992 | -104.19
9.7029 | -104.1922
9.7065 | -104.1948
9.7102 | -104.1984
9.7138 | -104.2065
9.7175 | -104.2169
9.7212 | -104.2246
9.7248 | -104.2294
9.7285 | -104.2336
9.7322 | -104.2387
9.7358 | -104.2456
9.7395 | -104.2539
9.7431 | -104.2621
9.7468 | -104.2692
9.7505 | -104.2797
9.7541 | -104.2944
9.7578 | -104.3066
9.7615 | -104.3188
9.7651 | -104.3331
9.7688 | -104.3555
9.7724 | -104.3756
9.7761 | -104.3928
9.7798 | -104.4087
9.7834 | -104.4281
9.7871 | -104.4498
9.7908 | -104.4731
9.7944 | -104.4948
9.7981 | -104.511
9.8017 | -104.5219
9.8054 | -104.5322
9.8091 | -104.5426
9.8127 | -104.5523
9.8164 | -104.55
9.8201 | -104.5422
9.8237 | -104.5306
9.8274 | -104.5135
9.8311 | -104.4938
9.8347 | -104.4727
9.8384 | -104.4491
9.842 | -104.4239
9.8457 | -104.3935
9.8494 | -104.3614
9.853 | -104.3246
9.8567 | -104.289
9.8604 | -104.2588
9.864 | -104.2326
9.8677 | -104.2054
9.8713 | -104.1811
9.875 | -104.1611
9.8787 | -104.1446
9.8823 | -104.1291
9.886 | -104.1178
9.8897 | -104.1123
9.8933 | -104.1119
9.897 | -104.1174
9.9006 | -104.1311
9.9043 | -104.1444
9.908 | -104.1575
9.9116 | -104.1741
9.9153 | -104.1968
9.919 | -104.2212
9.9226 | -104.2475
9.9263 | -104.2723
9.9299 | -104.2916
9.9336 | -104.313
9.9373 | -104.334
9.9409 | -104.3547
9.9446 | -104.3767
9.9483 | -104.398
9.9519 | -104.4124
9.9556 | -104.4233
9.9593 | -104.439
9.9629 | -104.4532
9.9666 | -104.4636
9.9702 | -104.4734
9.9739 | -104.485
9.9776 | -104.4938
9.9812 | -104.5007
9.9849 | -104.5077
9.9886 | -104.5176
9.9922 | -104.5259
9.9959 | -104.5335
9.9995 | -104.5387
10.0032 | -104.5371
10.0069 | -104.5362
10.0105 | -104.5398
10.0142 | -104.5395
10.0179 | -104.5338
10.0215 | -104.5285
10.0252 | -104.5247
10.0288 | -104.5238
10.0325 | -104.5201
10.0362 | -104.5135
10.0398 | -104.5227
10.0435 | -104.5202
10.0472 | -104.5262
10.0508 | -104.5401
10.0545 | -104.5554
10.0581 | -104.5698
10.0618 | -104.5891
10.0655 | -104.6101
10.0691 | -104.6326
10.0728 | -104.6541
10.0765 | -104.6804
10.0801 | -104.7042
10.0838 | -104.7307
10.0875 | -104.7582
10.0911 | -104.7785
10.0948 | -104.7927
10.0984 | -104.8089
10.1021 | -104.8216
10.1058 | -104.8288
10.1094 | -104.8304
10.1131 | -104.8245
10.1168 | -104.8135
10.1204 | -104.7995
10.1241 | -104.7848
10.1277 | -104.7696
10.1314 | -104.7484
10.1351 | -104.7263
10.1387 | -104.7021
10.1424 | -104.674
10.1461 | -104.6449
10.1497 | -104.6225
10.1534 | -104.5944
10.157 | -104.5629
10.1607 | -104.5401
10.1644 | -104.5227
10.168 | -104.5038
10.1717 | -104.4835
10.1754 | -104.4604
10.179 | -104.4408
10.1827 | -104.421
10.1863 | -104.4005
10.19 | -104.3809
10.1937 | -104.3651
10.1973 | -104.3444
10.201 | -104.3255
10.2047 | -104.3082
10.2083 | -104.2881
10.212 | -104.2671
10.2156 | -104.2495
10.2193 | -104.232
10.223 | -104.2159
10.2266 | -104.1989
10.2303 | -104.184
10.234 | -104.1696
10.2376 | -104.156
10.2413 | -104.1514
10.245 | -104.1495
10.2486 | -104.1508
10.2523 | -104.1575
10.2559 | -104.1679
10.2596 | -104.1898
10.2633 | -104.2176
10.2669 | -104.2529
10.2706 | -104.2918
10.2743 | -104.334
10.2779 | -104.3785
10.2816 | -104.4315
10.2852 | -104.4858
10.2889 | -104.5401
10.2926 | -104.5967
10.2962 | -104.6531
10.2999 | -104.7072
10.3036 | -104.7649
10.3072 | -104.8209
10.3109 | -104.8685
10.3145 | -104.9045
10.3182 | -104.9387
10.3219 | -104.9708
10.3255 | -104.9968
10.3292 | -105.0156
10.3329 | -105.0289
10.3365 | -105.032
10.3402 | -105.0291
10.3438 | -105.0151
10.3475 | -104.9964
10.3512 | -104.9756
10.3548 | -104.9548
10.3585 | -104.9293
10.3622 | -104.8988
10.3658 | -104.8662
10.3695 | -104.8276
10.3732 | -104.788
10.3768 | -104.7538
10.3805 | -104.721
10.3841 | -104.6888
10.3878 | -104.6561
10.3915 | -104.6244
10.3951 | -104.5993
10.3988 | -104.5796
10.4025 | -104.5637
10.4061 | -104.552
10.4098 | -104.5403
10.4134 | -104.537
10.4171 | -104.5347
10.4208 | -104.5378
10.4244 | -104.5442
10.4281 | -104.5536
10.4318 | -104.5661
10.4354 | -104.5813
10.4391 | -104.5975
10.4427 | -104.6131
10.4464 | -104.6255
10.4501 | -104.6323
10.4537 | -104.6374
10.4574 | -104.6456
10.4611 | -104.6534
10.4647 | -104.6661
10.4684 | -104.6762
10.472 | -104.6847
10.4757 | -104.6943
10.4794 | -104.7065
10.483 | -104.7195
10.4867 | -104.7371
10.4904 | -104.7574
10.494 | -104.7788
10.4977 | -104.8039
10.5014 | -104.8273
10.505 | -104.8512
10.5087 | -104.8822
10.5123 | -104.9138
10.516 | -104.9425
10.5197 | -104.9769
10.5233 | -105.0099
10.527 | -105.0351
10.5307 | -105.0563
10.5343 | -105.0743
10.538 | -105.0893
10.5416 | -105.0967
10.5453 | -105.0944
10.549 | -105.0836
10.5526 | -105.0722
10.5563 | -105.058
10.56 | -105.042
10.5636 | -105.0199
10.5673 | -104.9903
10.5709 | -104.9638
10.5746 | -104.9374
10.5783 | -104.9121
10.5819 | -104.8875
10.5856 | -104.8633
10.5893 | -104.839
10.5929 | -104.8178
10.5966 | -104.7973
10.6002 | -104.7784
10.6039 | -104.7617
10.6076 | -104.7438
10.6112 | -104.7266
10.6149 | -104.7136
10.6186 | -104.6984
10.6222 | -104.6871
10.6259 | -104.675
10.6296 | -104.6639
10.6332 | -104.6523
10.6369 | -104.6434
10.6405 | -104.6342
10.6442 | -104.6252
10.6479 | -104.6152
10.6515 | -104.5995
10.6552 | -104.5843
10.6589 | -104.572
10.6625 | -104.5579
10.6662 | -104.5453
10.6698 | -104.5332
10.6735 | -104.5197
10.6772 | -104.5085
10.6808 | -104.5029
10.6845 | -104.5014
10.6882 | -104.5023
10.6918 | -104.5083
10.6955 | -104.5153
10.6991 | -104.5239
10.7028 | -104.5328
10.7065 | -104.5414
10.7101 | -104.5511
10.7138 | -104.5622
10.7175 | -104.5718
10.7211 | -104.5807
10.7248 | -104.5913
10.7284 | -104.6035
10.7321 | -104.6118
10.7358 | -104.6199
10.7394 | -104.6237
10.7431 | -104.6208
10.7468 | -104.6182
10.7504 | -104.6171
10.7541 | -104.6158
10.7577 | -104.612
10.7614 | -104.6078
10.7651 | -104.6051
10.7687 | -104.5966
10.7724 | -104.5886
10.7761 | -104.5853
10.7797 | -104.5819
10.7834 | -104.5804
10.7871 | -104.578
10.7907 | -104.5757
10.7944 | -104.5685
10.798 | -104.5641
10.8017 | -104.5647
10.8054 | -104.5648
10.809 | -104.5665
10.8127 | -104.5684
10.8164 | -104.5671
10.82 | -104.5691
10.8237 | -104.5698
10.8273 | -104.5696
10.831 | -104.5683
10.8347 | -104.5674
10.8383 | -104.5647
10.842 | -104.5587
10.8457 | -104.5522
10.8493 | -104.5494
10.853 | -104.553
10.8566 | -104.5583
10.8603 | -104.5681
10.864 | -104.5759
10.8676 | -104.5874
10.8713 | -104.6041
10.875 | -104.626
10.8786 | -104.6485
10.8823 | -104.6737
10.8859 | -104.702
10.8896 | -104.7328
10.8933 | -104.7642
10.8969 | -104.7963
10.9006 | -104.8346
10.9043 | -104.8727
10.9079 | -104.9036
10.9116 | -104.9339
10.9153 | -104.9637
10.9189 | -104.9848
10.9226 | -104.9986
10.9262 | -105.0098
10.9299 | -105.0153
10.9336 | -105.0151
10.9372 | -105.0087
10.9409 | -104.9978
10.9446 | -104.9858
10.9482 | -104.9716
10.9519 | -104.9564
10.9555 | -104.9346
10.9592 | -104.9046
10.9629 | -104.8778
10.9665 | -104.852
10.9702 | -104.8232
10.9739 | -104.7954
10.9775 | -104.7694
10.9812 | -104.7468
10.9848 | -104.7283
10.9885 | -104.7147
10.9922 | -104.6957
10.9958 | -104.6818
10.9995 | -104.6765
11.0032 | -104.6728
11.0068 | -104.6711
11.0105 | -104.6722
11.0141 | -104.6784
11.0178 | -104.6842
11.0215 | -104.6963
11.0251 | -104.7094
11.0288 | -104.7163
11.0325 | -104.7264
11.0361 | -104.7411
11.0398 | -104.7548
11.0434 | -104.7676
11.0471 | -104.7816
11.0508 | -104.7946
11.0544 | -104.8092
11.0581 | -104.8205
11.0618 | -104.8314
11.0654 | -104.8463
11.0691 | -104.8571
11.0728 | -104.8699
11.0764 | -104.8844
11.0801 | -104.8907
11.0837 | -104.897
11.0874 | -104.9025
11.0911 | -104.9071
11.0947 | -104.9062
11.0984 | -104.9072
11.1021 | -104.9088
11.1057 | -104.9077
11.1094 | -104.9006
11.113 | -104.8947
11.1167 | -104.892
11.1204 | -104.8939
11.124 | -104.8915
11.1277 | -104.888
11.1314 | -104.888
11.135 | -104.8898
11.1387 | -104.8881
11.1423 | -104.8908
11.146 | -104.8955
11.1497 | -104.9004
11.1533 | -104.9153
11.157 | -104.9261
11.1607 | -104.9348
11.1643 | -104.9452
11.168 | -104.9504
11.1716 | -104.9583
11.1753 | -104.9667
11.179 | -104.9744
11.1826 | -104.9807
11.1863 | -104.9888
11.19 | -104.999
11.1936 | -105.0047
11.1973 | -105.0056
11.201 | -105.004
11.2046 | -105.0008
11.2083 | -104.9936
11.2119 | -104.9838
11.2156 | -104.9729
11.2193 | -104.9598
11.2229 | -104.9452
11.2266 | -104.9286
11.2303 | -104.9138
11.2339 | -104.8949
11.2376 | -104.8715
11.2412 | -104.8541
11.2449 | -104.8311
11.2486 | -104.8038
11.2522 | -104.7799
11.2559 | -104.7537
11.2596 | -104.731
11.2632 | -104.7131
11.2669 | -104.695
11.2705 | -104.6769
11.2742 | -104.6637
11.2779 | -104.6557
11.2815 | -104.6465
11.2852 | -104.6392
11.2889 | -104.6401
11.2925 | -104.6455
11.2962 | -104.6537
11.2998 | -104.6665
11.3035 | -104.6836
11.3072 | -104.7004
11.3108 | -104.7215
11.3145 | -104.7447
11.3182 | -104.7659
11.3218 | -104.7878
11.3255 | -104.8115
11.3292 | -104.8361
11.3328 | -104.8604
11.3365 | -104.8794
11.3401 | -104.8957
11.3438 | -104.9112
11.3475 | -104.9238
11.3511 | -104.9306
11.3548 | -104.9338
11.3585 | -104.9341
11.3621 | -104.9339
11.3658 | -104.9332
11.3694 | -104.9301
11.3731 | -104.9296
11.3768 | -104.9262
11.3804 | -104.9178
11.3841 | -104.9081
11.3878 | -104.8998
11.3914 | -104.8932
11.3951 | -104.8896
11.3987 | -104.8888
11.4024 | -104.8942
11.4061 | -104.9003
11.4097 | -104.9097
11.4134 | -104.9243
11.4171 | -104.9384
11.4207 | -104.9612
11.4244 | -104.9837
11.428 | -105.0057
11.4317 | -105.0328
11.4354 | -105.0647
11.439 | -105.1004
11.4427 | -105.1331
11.4464 | -105.1703
11.45 | -105.2077
11.4537 | -105.2429
11.4574 | -105.2762
11.461 | -105.3074
11.4647 | -105.3357
11.4683 | -105.3639
11.472 | -105.389
11.4757 | -105.4116
11.4793 | -105.4321
11.483 | -105.4482
11.4867 | -105.4636
11.4903 | -105.4761
11.494 | -105.4833
11.4976 | -105.4889
11.5013 | -105.488
11.505 | -105.4852
11.5086 | -105.4838
11.5123 | -105.4822
11.516 | -105.4775
11.5196 | -105.4712
11.5233 | -105.4653
11.5269 | -105.4571
11.5306 | -105.4453
11.5343 | -105.4352
11.5379 | -105.421
11.5416 | -105.4071
11.5453 | -105.3968
11.5489 | -105.3864
11.5526 | -105.3731
11.5562 | -105.3607
11.5599 | -105.3473
11.5636 | -105.3364
11.5672 | -105.3281
11.5709 | -105.3216
11.5746 | -105.3194
11.5782 | -105.3208
11.5819 | -105.3233
11.5855 | -105.3272
11.5892 | -105.3374
11.5929 | -105.3522
11.5965 | -105.3659
11.6002 | -105.3774
11.6039 | -105.3899
11.6075 | -105.4035
11.6112 | -105.4157
11.6149 | -105.4272
11.6185 | -105.437
11.6222 | -105.4353
11.6258 | -105.4412
11.6295 | -105.447
11.6332 | -105.4497
11.6368 | -105.4496
11.6405 | -105.4485
11.6442 | -105.4403
11.6478 | -105.4338
11.6515 | -105.4306
11.6551 | -105.4242
11.6588 | -105.4112
11.6625 | -105.4001
11.6661 | -105.391
11.6698 | -105.3845
11.6735 | -105.3808
11.6771 | -105.3687
11.6808 | -105.3559
11.6844 | -105.346
11.6881 | -105.3368
11.6918 | -105.3287
11.6954 | -105.3182
11.6991 | -105.3113
11.7028 | -105.3048
11.7064 | -105.3016
11.7101 | -105.3006
11.7137 | -105.2943
11.7174 | -105.2895
11.7211 | -105.2859
11.7247 | -105.2799
11.7284 | -105.2748
11.7321 | -105.2708
11.7357 | -105.263
11.7394 | -105.2568
11.7431 | -105.2491
11.7467 | -105.2432
11.7504 | -105.2422
11.754 | -105.2394
11.7577 | -105.237
11.7614 | -105.2343
11.765 | -105.2334
11.7687 | -105.235
11.7724 | -105.2347
11.776 | -105.2339
11.7797 | -105.2286
11.7833 | -105.2225
11.787 | -105.2214
11.7907 | -105.2215
11.7943 | -105.2184
11.798 | -105.2133
11.8017 | -105.2071
11.8053 | -105.2031
11.809 | -105.2009
11.8126 | -105.1973
11.8163 | -105.1997
11.82 | -105.1971
11.8236 | -105.1899
11.8273 | -105.183
11.831 | -105.1826
11.8346 | -105.1835
11.8383 | -105.181
11.8419 | -105.1802
11.8456 | -105.1759
11.8493 | -105.176
11.8529 | -105.1775
11.8566 | -105.1768
11.8603 | -105.1757
11.8639 | -105.1712
11.8676 | -105.1667
11.8713 | -105.1618
11.8749 | -105.1581
11.8786 | -105.1512
11.8822 | -105.1466
11.8859 | -105.1489
11.8896 | -105.1489
11.8932 | -105.1497
11.8969 | -105.1544
11.9006 | -105.1595
11.9042 | -105.1653
11.9079 | -105.1765
11.9115 | -105.1892
11.9152 | -105.2037
11.9189 | -105.22
11.9225 | -105.2366
11.9262 | -105.254
11.9299 | -105.2724
11.9335 | -105.2943
11.9372 | -105.317
11.9408 | -105.3396
11.9445 | -105.3627
11.9482 | -105.3868
11.9518 | -105.4148
11.9555 | -105.4414
11.9592 | -105.4593
11.9628 | -105.476
11.9665 | -105.4934
11.9701 | -105.5058
11.9738 | -105.5161
11.9775 | -105.5243
11.9811 | -105.5262
11.9848 | -105.5239
11.9885 | -105.5156
11.9921 | -105.5086
11.9958 | -105.5054
11.9995 | -105.5001
12.0031 | -105.4879
12.0068 | -105.4739
12.0104 | -105.4606
12.0141 | -105.4419
12.0178 | -105.4248
12.0214 | -105.4073
12.0251 | -105.3885
12.0288 | -105.3713
12.0324 | -105.3556
12.0361 | -105.3407
12.0397 | -105.325
12.0434 | -105.3043
12.0471 | -105.2842
12.0507 | -105.2696
12.0544 | -105.2545
12.0581 | -105.2385
12.0617 | -105.2266
12.0654 | -105.2152
12.069 | -105.2041
12.0727 | -105.1961
12.0764 | -105.1896
12.08 | -105.1814
12.0837 | -105.177
12.0874 | -105.1753
12.091 | -105.1735
12.0947 | -105.1729
12.0983 | -105.1712
12.102 | -105.1721
12.1057 | -105.1705
12.1093 | -105.1701
12.113 | -105.1731
12.1167 | -105.176
12.1203 | -105.1787
12.124 | -105.182
12.1276 | -105.1865
12.1313 | -105.194
12.135 | -105.2007
12.1386 | -105.21
12.1423 | -105.2219
12.146 | -105.2333
12.1496 | -105.2438
12.1533 | -105.2592
12.157 | -105.2706
12.1606 | -105.2868
12.1643 | -105.3004
12.1679 | -105.3107
12.1716 | -105.3209
12.1753 | -105.3318
12.1789 | -105.3425
12.1826 | -105.3502
12.1863 | -105.3566
12.1899 | -105.3636
12.1936 | -105.3683
12.1972 | -105.3688
12.2009 | -105.3677
12.2046 | -105.366
12.2082 | -105.361
12.2119 | -105.3548
12.2156 | -105.3505
12.2192 | -105.3468
12.2229 | -105.3414
12.2265 | -105.3312
12.2302 | -105.3186
12.2339 | -105.3095
12.2375 | -105.3016
12.2412 | -105.2958
12.2449 | -105.2932
12.2485 | -105.2903
12.2522 | -105.2878
12.2558 | -105.2895
12.2595 | -105.288
12.2632 | -105.2875
12.2668 | -105.2893
12.2705 | -105.2909
12.2742 | -105.2961
12.2778 | -105.3004
12.2815 | -105.3072
12.2852 | -105.3118
12.2888 | -105.3141
12.2925 | -105.3166
12.2961 | -105.3183
12.2998 | -105.3192
12.3035 | -105.3204
12.3071 | -105.3229
12.3108 | -105.3232
12.3145 | -105.3249
12.3181 | -105.3265
12.3218 | -105.3267
12.3254 | -105.3236
12.3291 | -105.3198
12.3328 | -105.3179
12.3364 | -105.3189
12.3401 | -105.3156
12.3438 | -105.3138
12.3474 | -105.3119
12.3511 | -105.3038
12.3547 | -105.2944
12.3584 | -105.2882
12.3621 | -105.283
12.3657 | -105.2772
12.3694 | -105.2736
12.3731 | -105.2691
12.3767 | -105.2648
12.3804 | -105.2616
12.384 | -105.2597
12.3877 | -105.2576
12.3914 | -105.2579
12.395 | -105.2596
12.3987 | -105.2621
12.4024 | -105.2618
12.406 | -105.2706
12.4097 | -105.285
12.4134 | -105.2964
12.417 | -105.3075
12.4207 | -105.3207
12.4243 | -105.3358
12.428 | -105.3586
12.4317 | -105.3808
12.4353 | -105.3973
12.439 | -105.4128
12.4427 | -105.4286
12.4463 | -105.4458
12.45 | -105.4612
12.4536 | -105.474
12.4573 | -105.4812
12.461 | -105.4829
12.4646 | -105.4856
12.4683 | -105.482
12.472 | -105.4762
12.4756 | -105.4676
12.4793 | -105.4563
12.4829 | -105.4442
12.4866 | -105.429
12.4903 | -105.413
12.4939 | -105.3944
12.4976 | -105.3769
12.5013 | -105.3595
12.5049 | -105.3434
12.5086 | -105.3269
12.5122 | -105.3122
12.5159 | -105.2974
12.5196 | -105.287
12.5232 | -105.2792
12.5269 | -105.2763
12.5306 | -105.2769
12.5342 | -105.28
12.5379 | -105.2871
12.5416 | -105.2959
12.5452 | -105.311
12.5489 | -105.3284
12.5525 | -105.3502
12.5562 | -105.3779
12.5599 | -105.4067
12.5635 | -105.4356
12.5672 | -105.4665
12.5709 | -105.4993
12.5745 | -105.5328
12.5782 | -105.5686
12.5818 | -105.6067
12.5855 | -105.6378
12.5892 | -105.6636
12.5928 | -105.69
12.5965 | -105.7149
12.6002 | -105.7356
12.6038 | -105.7512
12.6075 | -105.7589
12.6111 | -105.768
12.6148 | -105.7791
12.6185 | -105.7843
12.6221 | -105.7842
12.6258 | -105.778
12.6295 | -105.7695
12.6331 | -105.7599
12.6368 | -105.7488
12.6404 | -105.7375
12.6441 | -105.7239
12.6478 | -105.7137
12.6514 | -105.6993
12.6551 | -105.6857
12.6588 | -105.6731
12.6624 | -105.6589
12.6661 | -105.6449
12.6698 | -105.634
12.6734 | -105.6249
12.6771 | -105.6145
12.6807 | -105.6111
12.6844 | -105.6073
12.6881 | -105.6047
12.6917 | -105.6008
12.6954 | -105.597
12.6991 | -105.5956
12.7027 | -105.5948
12.7064 | -105.5949
12.71 | -105.5939
12.7137 | -105.5986
12.7174 | -105.6019
12.721 | -105.6027
12.7247 | -105.6067
12.7284 | -105.6086
12.732 | -105.6124
12.7357 | -105.6156
12.7393 | -105.6183
12.743 | -105.6231
12.7467 | -105.6281
12.7503 | -105.6341
12.754 | -105.6402
12.7577 | -105.6476
12.7613 | -105.655
12.765 | -105.6618
12.7686 | -105.6703
12.7723 | -105.6818
12.776 | -105.6936
12.7796 | -105.7046
12.7833 | -105.7183
12.787 | -105.7326
12.7906 | -105.745
12.7943 | -105.7572
12.7979 | -105.7709
12.8016 | -105.7878
12.8053 | -105.8004
12.8089 | -105.8134
12.8126 | -105.8267
12.8163 | -105.8371
12.8199 | -105.8459
12.8236 | -105.8541
12.8273 | -105.863
12.8309 | -105.8702
12.8346 | -105.8763
12.8382 | -105.8816
12.8419 | -105.8827
12.8456 | -105.8852
12.8492 | -105.8889
12.8529 | -105.8915
12.8566 | -105.8923
12.8602 | -105.8935
12.8639 | -105.8945
12.8675 | -105.889
12.8712 | -105.8807
12.8749 | -105.8743
12.8785 | -105.8719
12.8822 | -105.868
12.8859 | -105.8614
12.8895 | -105.8545
12.8932 | -105.8485
12.8968 | -105.8424
12.9005 | -105.8315
12.9042 | -105.8199
12.9078 | -105.8085
12.9115 | -105.7985
12.9152 | -105.7891
12.9188 | -105.7769
12.9225 | -105.7657
12.9261 | -105.7563
12.9298 | -105.7447
12.9335 | -105.7382
12.9371 | -105.7297
12.9408 | -105.7219
12.9445 | -105.7164
12.9481 | -105.7099
12.9518 | -105.7049
12.9555 | -105.7016
12.9591 | -105.6978
12.9628 | -105.6971
12.9664 | -105.6976
12.9701 | -105.6987
12.9738 | -105.6973
12.9774 | -105.6953
12.9811 | -105.693
12.9848 | -105.6869
12.9884 | -105.6811
12.9921 | -105.6775
12.9957 | -105.6694
12.9994 | -105.6617
13.0031 | -105.6543
13.0067 | -105.649
13.0104 | -105.6428
13.0141 | -105.6328
13.0177 | -105.623
13.0214 | -105.6146
13.025 | -105.6048
13.0287 | -105.597
13.0324 | -105.5887
13.036 | -105.5806
13.0397 | -105.5713
13.0434 | -105.5626
13.047 | -105.5574
13.0507 | -105.5542
13.0543 | -105.5498
13.058 | -105.5463
13.0617 | -105.5448
13.0653 | -105.5413
13.069 | -105.5393
13.0727 | -105.5345
13.0763 | -105.5272
13.08 | -105.5217
13.0837 | -105.5149
13.0873 | -105.5116
13.091 | -105.5081
13.0946 | -105.5013
13.0983 | -105.4889
13.102 | -105.4745
13.1056 | -105.4592
13.1093 | -105.4414
13.113 | -105.422
13.1166 | -105.4022
13.1203 | -105.3812
13.1239 | -105.3627
13.1276 | -105.3419
13.1313 | -105.3182
13.1349 | -105.2929
13.1386 | -105.2702
13.1423 | -105.2518
13.1459 | -105.2358
13.1496 | -105.2166
13.1532 | -105.2012
13.1569 | -105.1873
13.1606 | -105.177
13.1642 | -105.1719
13.1679 | -105.1686
13.1716 | -105.1677
13.1752 | -105.1642
13.1789 | -105.1635
13.1825 | -105.1667
13.1862 | -105.1696
13.1899 | -105.1763
13.1935 | -105.1841
13.1972 | -105.1931
13.2009 | -105.2035
13.2045 | -105.2154
13.2082 | -105.228
13.2118 | -105.2419
13.2155 | -105.2566
13.2192 | -105.2719
13.2228 | -105.286
13.2265 | -105.3013
13.2302 | -105.3176
13.2338 | -105.3343
13.2375 | -105.3498
13.2412 | -105.3665
13.2448 | -105.3854
13.2485 | -105.4043
13.2521 | -105.4205
13.2558 | -105.4377
13.2595 | -105.4533
13.2631 | -105.4711
13.2668 | -105.4945
13.2705 | -105.5144
13.2741 | -105.5353
13.2778 | -105.5544
13.2814 | -105.5743
13.2851 | -105.5947
13.2888 | -105.6169
13.2924 | -105.6427
13.2961 | -105.6712
13.2998 | -105.6973
13.3034 | -105.7217
13.3071 | -105.7436
13.3107 | -105.7614
13.3144 | -105.7824
13.3181 | -105.8034
13.3217 | -105.8219
13.3254 | -105.841
13.3291 | -105.8564
13.3327 | -105.8642
13.3364 | -105.8675
13.34 | -105.8662
13.3437 | -105.8627
13.3474 | -105.8577
13.351 | -105.8494
13.3547 | -105.8388
13.3584 | -105.8227
13.362 | -105.802
13.3657 | -105.7849
13.3694 | -105.7638
13.373 | -105.7377
13.3767 | -105.7115
13.3803 | -105.6825
13.384 | -105.6517
13.3877 | -105.6214
13.3913 | -105.5925
13.395 | -105.5672
13.3987 | -105.5436
13.4023 | -105.5218
13.406 | -105.4986
13.4096 | -105.4783
13.4133 | -105.459
13.417 | -105.4467
13.4206 | -105.4369
13.4243 | -105.4297
13.428 | -105.4245
13.4316 | -105.4232
13.4353 | -105.4246
13.4389 | -105.4272
13.4426 | -105.4316
13.4463 | -105.4405
13.4499 | -105.4513
13.4536 | -105.465
13.4573 | -105.4821
13.4609 | -105.4936
13.4646 | -105.5074
13.4682 | -105.5227
13.4719 | -105.5393
13.4756 | -105.5562
13.4792 | -105.5723
13.4829 | -105.5902
13.4866 | -105.6071
13.4902 | -105.6195
13.4939 | -105.6333
13.4976 | -105.6484
13.5012 | -105.6672
13.5049 | -105.6856
13.5085 | -105.6999
13.5122 | -105.7146
13.5159 | -105.7298
13.5195 | -105.746
13.5232 | -105.7581
13.5269 | -105.7704
13.5305 | -105.7827
13.5342 | -105.7931
13.5378 | -105.8008
13.5415 | -105.8032
13.5452 | -105.8046
13.5488 | -105.8019
13.5525 | -105.8007
13.5562 | -105.7993
13.5598 | -105.7923
13.5635 | -105.7857
13.5671 | -105.7807
13.5708 | -105.7732
13.5745 | -105.7622
13.5781 | -105.7529
13.5818 | -105.7442
13.5855 | -105.734
13.5891 | -105.7205
13.5928 | -105.7064
13.5964 | -105.6929
13.6001 | -105.6774
13.6038 | -105.6621
13.6074 | -105.6518
13.6111 | -105.6387
13.6148 | -105.6238
13.6184 | -105.6107
13.6221 | -105.5968
13.6258 | -105.5813
13.6294 | -105.5658
13.6331 | -105.5527
13.6367 | -105.5396
13.6404 | -105.5285
13.6441 | -105.5176
13.6477 | -105.5059
13.6514 | -105.4952
13.6551 | -105.4894
13.6587 | -105.4849
13.6624 | -105.4804
13.666 | -105.4751
13.6697 | -105.4725
13.6734 | -105.4743
13.677 | -105.4738
13.6807 | -105.4738
13.6844 | -105.4788
13.688 | -105.4875
13.6917 | -105.4953
13.6953 | -105.5031
13.699 | -105.5133
13.7027 | -105.5263
13.7063 | -105.5411
13.71 | -105.5577
13.7137 | -105.5729
13.7173 | -105.5908
13.721 | -105.6111
13.7246 | -105.631
13.7283 | -105.6555
13.732 | -105.6779
13.7356 | -105.6977
13.7393 | -105.7207
13.743 | -105.7443
13.7466 | -105.7692
13.7503 | -105.7901
13.7539 | -105.8093
13.7576 | -105.829
13.7613 | -105.8471
13.7649 | -105.8638
13.7686 | -105.8776
13.7723 | -105.8936
13.7759 | -105.9058
13.7796 | -105.9174
13.7833 | -105.9285
13.7869 | -105.9374
13.7906 | -105.9446
13.7942 | -105.9467
13.7979 | -105.9512
13.8016 | -105.9538
13.8052 | -105.9525
13.8089 | -105.9517
13.8126 | -105.952
13.8162 | -105.9493
13.8199 | -105.9463
13.8235 | -105.9479
13.8272 | -105.9487
13.8309 | -105.9476
13.8345 | -105.947
13.8382 | -105.9446
13.8419 | -105.9456
13.8455 | -105.9505
13.8492 | -105.9571
13.8528 | -105.9619
13.8565 | -105.9662
13.8602 | -105.971
13.8638 | -105.977
13.8675 | -105.9822
13.8712 | -105.9835
13.8748 | -105.9857
13.8785 | -105.9895
13.8821 | -105.9926
13.8858 | -105.9933
13.8895 | -105.9945
13.8931 | -105.9934
13.8968 | -105.9913
13.9005 | -105.9875
13.9041 | -105.9821
13.9078 | -105.9783
13.9115 | -105.9744
13.9151 | -105.9683
13.9188 | -105.9595
13.9224 | -105.9507
13.9261 | -105.9421
13.9298 | -105.9327
13.9334 | -105.9219
13.9371 | -105.9116
13.9408 | -105.8992
13.9444 | -105.8864
13.9481 | -105.8734
13.9517 | -105.8634
13.9554 | -105.8515
13.9591 | -105.8392
13.9627 | -105.8301
13.9664 | -105.8198
13.9701 | -105.8129
13.9737 | -105.8077
13.9774 | -105.8042
13.981 | -105.8017
13.9847 | -105.7995
13.9884 | -105.7971
13.992 | -105.7978
13.9957 | -105.8046
13.9994 | -105.8084
14.003 | -105.812
14.0067 | -105.8161
14.0103 | -105.8216
14.014 | -105.8297
14.0177 | -105.8394
14.0213 | -105.8502
14.025 | -105.8557
14.0287 | -105.8599
14.0323 | -105.8638
14.036 | -105.8665
14.0396 | -105.8677
14.0433 | -105.8646
14.047 | -105.858
14.0506 | -105.8464
14.0543 | -105.8379
14.058 | -105.8286
14.0616 | -105.8194
14.0653 | -105.8116
14.069 | -105.8027
14.0726 | -105.7955
14.0763 | -105.7893
14.0799 | -105.7786
14.0836 | -105.7669
14.0873 | -105.7563
14.0909 | -105.7469
14.0946 | -105.7396
14.0983 | -105.7348
14.1019 | -105.7317
14.1056 | -105.7308
14.1092 | -105.7305
14.1129 | -105.7324
14.1166 | -105.7366
14.1202 | -105.7436
14.1239 | -105.7518
14.1276 | -105.7598
14.1312 | -105.771
14.1349 | -105.785
14.1385 | -105.8018
14.1422 | -105.8197
14.1459 | -105.8398
14.1495 | -105.8596
14.1532 | -105.8808
14.1569 | -105.906
14.1605 | -105.931
14.1642 | -105.9532
14.1678 | -105.9778
14.1715 | -106.0031
14.1752 | -106.0268
14.1788 | -106.0505
14.1825 | -106.0719
14.1862 | -106.0941
14.1898 | -106.1125
14.1935 | -106.1318
14.1972 | -106.1498
14.2008 | -106.166
14.2045 | -106.1788
14.2081 | -106.189
14.2118 | -106.1941
14.2155 | -106.1972
14.2191 | -106.1977
14.2228 | -106.1963
14.2265 | -106.1941
14.2301 | -106.1881
14.2338 | -106.1803
14.2374 | -106.1732
14.2411 | -106.1638
14.2448 | -106.1532
14.2484 | -106.1441
14.2521 | -106.1352
14.2558 | -106.1204
14.2594 | -106.1066
14.2631 | -106.094
14.2667 | -106.0836
14.2704 | -106.0692
14.2741 | -106.0559
14.2777 | -106.0411
14.2814 | -106.0279
14.2851 | -106.0161
14.2887 | -106.0075
14.2924 | -106.0013
14.296 | -105.9941
14.2997 | -105.989
14.3034 | -105.9847
14.307 | -105.9796
14.3107 | -105.9751
14.3144 | -105.97
14.318 | -105.9672
14.3217 | -105.963
14.3254 | -105.9606
14.329 | -105.9605
14.3327 | -105.9568
14.3363 | -105.9531
14.34 | -105.9493
14.3437 | -105.9424
14.3473 | -105.9359
14.351 | -105.931
14.3547 | -105.9281
14.3583 | -105.9245
14.362 | -105.9228
14.3656 | -105.9188
14.3693 | -105.9147
14.373 | -105.9106
14.3766 | -105.9064
14.3803 | -105.9024
14.384 | -105.8996
14.3876 | -105.9011
14.3913 | -105.905
14.3949 | -105.9106
14.3986 | -105.9126
14.4023 | -105.9143
14.4059 | -105.9181
14.4096 | -105.9245
14.4133 | -105.9332
14.4169 | -105.9433
14.4206 | -105.9539
14.4242 | -105.9651
14.4279 | -105.9755
14.4316 | -105.9873
14.4352 | -106.0035
14.4389 | -106.0192
14.4426 | -106.0321
14.4462 | -106.0417
14.4499 | -106.0547
14.4536 | -106.0669
14.4572 | -106.0788
14.4609 | -106.0925
14.4645 | -106.1046
14.4682 | -106.1116
14.4719 | -106.1175
14.4755 | -106.1214
14.4792 | -106.126
14.4829 | -106.1294
14.4865 | -106.1358
14.4902 | -106.1378
14.4938 | -106.139
14.4975 | -106.1373
14.5012 | -106.1312
14.5048 | -106.1227
14.5085 | -106.1121
14.5122 | -106.1014
14.5158 | -106.0926
14.5195 | -106.0825
14.5231 | -106.0752
14.5268 | -106.0671
14.5305 | -106.0577
14.5341 | -106.0458
14.5378 | -106.0335
14.5415 | -106.0239
14.5451 | -106.0129
14.5488 | -106.0016
14.5524 | -105.9936
14.5561 | -105.9859
14.5598 | -105.9734
14.5634 | -105.9619
14.5671 | -105.9538
14.5708 | -105.9434
14.5744 | -105.9347
14.5781 | -105.9261
14.5817 | -105.9187
14.5854 | -105.9093
14.5891 | -105.8982
14.5927 | -105.8876
14.5964 | -105.8752
14.6001 | -105.861
14.6037 | -105.8518
14.6074 | -105.8418
14.6111 | -105.8311
14.6147 | -105.8235
14.6184 | -105.8173
14.622 | -105.8096
14.6257 | -105.8021
14.6294 | -105.796
14.633 | -105.7941
14.6367 | -105.7931
14.6404 | -105.7905
14.644 | -105.7889
14.6477 | -105.7893
14.6513 | -105.7899
14.655 | -105.7959
14.6587 | -105.7965
14.6623 | -105.7987
14.666 | -105.8052
14.6697 | -105.8097
14.6733 | -105.8123
14.677 | -105.8215
14.6806 | -105.833
14.6843 | -105.8447
14.688 | -105.855
14.6916 | -105.8654
14.6953 | -105.8717
14.699 | -105.8791
14.7026 | -105.8892
14.7063 | -105.8997
14.7099 | -105.9128
14.7136 | -105.9238
14.7173 | -105.9331
14.7209 | -105.9438
14.7246 | -105.9547
14.7283 | -105.9662
14.7319 | -105.9761
14.7356 | -105.9872
14.7393 | -106.0017
14.7429 | -106.0144
14.7466 | -106.0267
14.7502 | -106.0403
14.7539 | -106.0546
14.7576 | -106.0676
14.7612 | -106.0754
14.7649 | -106.0825
14.7686 | -106.0918
14.7722 | -106.1006
14.7759 | -106.106
14.7795 | -106.1107
14.7832 | -106.1145
14.7869 | -106.1184
14.7905 | -106.1183
14.7942 | -106.1097
14.7979 | -106.1077
14.8015 | -106.1065
14.8052 | -106.1043
14.8088 | -106.1016
14.8125 | -106.0978
14.8162 | -106.0895
14.8198 | -106.0813
14.8235 | -106.0747
14.8272 | -106.0665
14.8308 | -106.0588
14.8345 | -106.0538
14.8381 | -106.0504
14.8418 | -106.0446
14.8455 | -106.0395
14.8491 | -106.0377
14.8528 | -106.0372
14.8565 | -106.0388
14.8601 | -106.0436
14.8638 | -106.0506
14.8675 | -106.057
14.8711 | -106.0678
14.8748 | -106.0768
14.8784 | -106.087
14.8821 | -106.1014
14.8858 | -106.12
14.8894 | -106.1351
14.8931 | -106.1458
14.8968 | -106.1608
14.9004 | -106.1766
14.9041 | -106.1921
14.9077 | -106.2101
14.9114 | -106.2277
14.9151 | -106.2424
14.9187 | -106.2547
14.9224 | -106.2659
14.9261 | -106.2771
14.9297 | -106.2877
14.9334 | -106.2939
14.937 | -106.2966
14.9407 | -106.2976
14.9444 | -106.2957
14.948 | -106.2942
14.9517 | -106.2909
14.9554 | -106.2861
14.959 | -106.2806
14.9627 | -106.2729
14.9663 | -106.2628
14.97 | -106.2526
14.9737 | -106.2426
14.9773 | -106.2282
14.981 | -106.2136
14.9847 | -106.1982
14.9883 | -106.1818
14.992 | -106.1644
14.9956 | -106.1479
14.9993 | -106.1308
15.003 | -106.1184
15.0066 | -106.1078
15.0103 | -106.0946
15.014 | -106.0829
15.0176 | -106.0673
15.0213 | -106.0519
15.025 | -106.0356
15.0286 | -106.0221
15.0323 | -106.0115
15.0359 | -106.0033
15.0396 | -105.9949
15.0433 | -105.9868
15.0469 | -105.9788
15.0506 | -105.9682
15.0543 | -105.9608
15.0579 | -105.9548
15.0616 | -105.9497
15.0652 | -105.9443
15.0689 | -105.9416
15.0726 | -105.9401
15.0762 | -105.9377
15.0799 | -105.9368
15.0836 | -105.9351
15.0872 | -105.9373
15.0909 | -105.9421
15.0945 | -105.9429
15.0982 | -105.9435
15.1019 | -105.9443
15.1055 | -105.9471
15.1092 | -105.953
15.1129 | -105.9564
15.1165 | -105.9611
15.1202 | -105.9643
15.1238 | -105.9664
15.1275 | -105.9698
15.1312 | -105.9731
15.1348 | -105.9758
15.1385 | -105.9787
15.1422 | -105.9816
15.1458 | -105.9807
15.1495 | -105.9777
15.1532 | -105.9738
15.1568 | -105.9701
15.1605 | -105.9667
15.1641 | -105.9576
15.1678 | -105.9444
15.1715 | -105.9359
15.1751 | -105.927
15.1788 | -105.9175
15.1825 | -105.9077
15.1861 | -105.8963
15.1898 | -105.8866
15.1934 | -105.8782
15.1971 | -105.8671
15.2008 | -105.856
15.2044 | -105.8473
15.2081 | -105.8385
15.2118 | -105.8291
15.2154 | -105.8196
15.2191 | -105.8079
15.2227 | -105.7977
15.2264 | -105.792
15.2301 | -105.7899
15.2337 | -105.7917
15.2374 | -105.7932
15.2411 | -105.7977
15.2447 | -105.801
15.2484 | -105.8047
15.252 | -105.808
15.2557 | -105.8139
15.2594 | -105.821
15.263 | -105.8276
15.2667 | -105.8367
15.2704 | -105.8464
15.274 | -105.8572
15.2777 | -105.8696
15.2814 | -105.881
15.285 | -105.8929
15.2887 | -105.9033
15.2923 | -105.9115
15.296 | -105.9207
15.2997 | -105.9319
15.3033 | -105.9418
15.307 | -105.95
15.3107 | -105.9585
15.3143 | -105.9648
15.318 | -105.9675
15.3216 | -105.9605
15.3253 | -105.9619
15.329 | -105.9638
15.3326 | -105.9647
15.3363 | -105.9615
15.34 | -105.9598
15.3436 | -105.9544
15.3473 | -105.9488
15.3509 | -105.9425
15.3546 | -105.9357
15.3583 | -105.9254
15.3619 | -105.9142
15.3656 | -105.9019
15.3693 | -105.8921
15.3729 | -105.8814
15.3766 | -105.8681
15.3802 | -105.856
15.3839 | -105.844
15.3876 | -105.8348
15.3912 | -105.8226
15.3949 | -105.8083
15.3986 | -105.7961
15.4022 | -105.7844
15.4059 | -105.7737
15.4096 | -105.7622
15.4132 | -105.7517
15.4169 | -105.7453
15.4205 | -105.7382
15.4242 | -105.7297
15.4279 | -105.7241
15.4315 | -105.7217
15.4352 | -105.719
15.4389 | -105.7154
15.4425 | -105.7124
15.4462 | -105.7105
15.4498 | -105.7047
15.4535 | -105.7003
15.4572 | -105.6988
15.4608 | -105.6985
15.4645 | -105.6997
15.4682 | -105.6998
15.4718 | -105.6994
15.4755 | -105.6989
15.4791 | -105.6991
15.4828 | -105.7003
15.4865 | -105.7027
15.4901 | -105.7013
15.4938 | -105.7024
15.4975 | -105.7017
15.5011 | -105.7042
15.5048 | -105.7053
15.5084 | -105.7054
15.5121 | -105.7082
15.5158 | -105.7104
15.5194 | -105.71
15.5231 | -105.7107
15.5268 | -105.7123
15.5304 | -105.7125
15.5341 | -105.7147
15.5378 | -105.7202
15.5414 | -105.7248
15.5451 | -105.729
15.5487 | -105.7326
15.5524 | -105.7344
15.5561 | -105.7333
15.5597 | -105.7349
15.5634 | -105.738
15.5671 | -105.7415
15.5707 | -105.7413
15.5744 | -105.7441
15.578 | -105.7482
15.5817 | -105.7532
15.5854 | -105.7582
15.589 | -105.7645
15.5927 | -105.77
15.5964 | -105.7776
15.6 | -105.7849
15.6037 | -105.794
15.6073 | -105.8042
15.611 | -105.8123
15.6147 | -105.8234
15.6183 | -105.8379
15.622 | -105.8505
15.6257 | -105.8659
15.6293 | -105.8845
15.633 | -105.9025
15.6366 | -105.9185
15.6403 | -105.9319
15.644 | -105.9481
15.6476 | -105.9675
15.6513 | -105.985
15.655 | -106.0015
15.6586 | -106.018
15.6623 | -106.0346
15.6659 | -106.0511
15.6696 | -106.0646
15.6733 | -106.0801
15.6769 | -106.0934
15.6806 | -106.1049
15.6843 | -106.1131
15.6879 | -106.1195
15.6916 | -106.1268
15.6953 | -106.1339
15.6989 | -106.1393
15.7026 | -106.1444
15.7062 | -106.1502
15.7099 | -106.155
15.7136 | -106.1614
15.7172 | -106.164
15.7209 | -106.1652
15.7246 | -106.1672
15.7282 | -106.1717
15.7319 | -106.1781
15.7355 | -106.1833
15.7392 | -106.1889
15.7429 | -106.196
15.7465 | -106.2014
15.7502 | -106.2078
15.7539 | -106.2129
15.7575 | -106.2183
15.7612 | -106.2251
15.7648 | -106.2315
15.7685 | -106.2354
15.7722 | -106.2429
15.7758 | -106.2535
15.7795 | -106.265
15.7832 | -106.274
15.7868 | -106.2815
15.7905 | -106.2913
15.7941 | -106.3026
15.7978 | -106.3082
15.8015 | -106.3122
15.8051 | -106.3143
15.8088 | -106.316
15.8125 | -106.3175
15.8161 | -106.3167
15.8198 | -106.3122
15.8235 | -106.3082
15.8271 | -106.2986
15.8308 | -106.2916
15.8344 | -106.2849
15.8381 | -106.2708
15.8418 | -106.2587
15.8454 | -106.2465
15.8491 | -106.2306
15.8528 | -106.2114
15.8564 | -106.1942
15.8601 | -106.181
15.8637 | -106.1708
15.8674 | -106.1585
15.8711 | -106.1468
15.8747 | -106.1353
15.8784 | -106.1237
15.8821 | -106.1104
15.8857 | -106.0999
15.8894 | -106.0918
15.893 | -106.0817
15.8967 | -106.0731
15.9004 | -106.065
15.904 | -106.0593
15.9077 | -106.0516
15.9114 | -106.0464
15.915 | -106.0435
15.9187 | -106.043
15.9223 | -106.0441
15.926 | -106.0414
15.9297 | -106.0397
15.9333 | -106.0398
15.937 | -106.0409
15.9407 | -106.0423
15.9443 | -106.0448
15.948 | -106.0492
15.9517 | -106.054
15.9553 | -106.0587
15.959 | -106.0648
15.9626 | -106.0723
15.9663 | -106.0797
15.97 | -106.0879
15.9736 | -106.0947
15.9773 | -106.1017
15.981 | -106.1129
15.9846 | -106.1218
15.9883 | -106.1273
15.9919 | -106.1348
15.9956 | -106.1429
15.9993 | -106.147
16.0029 | -106.1513
16.0066 | -106.1577
16.0103 | -106.1619
16.0139 | -106.1643
16.0176 | -106.165
16.0212 | -106.1654
16.0249 | -106.1662
16.0286 | -106.1699
16.0322 | -106.1747
16.0359 | -106.1793
16.0396 | -106.1823
16.0432 | -106.1824
16.0469 | -106.1832
16.0505 | -106.1849
16.0542 | -106.1889
16.0579 | -106.1912
16.0615 | -106.1929
16.0652 | -106.1977
16.0689 | -106.1984
16.0725 | -106.1993
16.0762 | -106.2039
16.0799 | -106.2078
16.0835 | -106.2155
16.0872 | -106.2204
16.0908 | -106.2216
16.0945 | -106.2231
16.0982 | -106.2261
16.1018 | -106.2282
16.1055 | -106.23
16.1092 | -106.2355
16.1128 | -106.2424
16.1165 | -106.2488
16.1201 | -106.2541
16.1238 | -106.2585
16.1275 | -106.2607
16.1311 | -106.2626
16.1348 | -106.2609
16.1385 | -106.2613
16.1421 | -106.2611
16.1458 | -106.2578
16.1494 | -106.2533
16.1531 | -106.2481
16.1568 | -106.2439
16.1604 | -106.2394
16.1641 | -106.2342
16.1678 | -106.2255
16.1714 | -106.217
16.1751 | -106.2071
16.1787 | -106.197
16.1824 | -106.1844
16.1861 | -106.1691
16.1897 | -106.1533
16.1934 | -106.1363
16.1971 | -106.1188
16.2007 | -106.104
16.2044 | -106.0872
16.208 | -106.0711
16.2117 | -106.0546
16.2154 | -106.0375
16.219 | -106.0205
16.2227 | -106.0012
16.2264 | -105.9854
16.23 | -105.9718
16.2337 | -105.9571
16.2374 | -105.9415
16.241 | -105.9301
16.2447 | -105.9201
16.2483 | -105.9129
16.252 | -105.9076
16.2557 | -105.904
16.2593 | -105.9002
16.263 | -105.8987
16.2667 | -105.8999
16.2703 | -105.903
16.274 | -105.9082
16.2776 | -105.919
16.2813 | -105.9317
16.285 | -105.9426
16.2886 | -105.9541
16.2923 | -105.9691
16.296 | -105.9855
16.2996 | -106.0045
16.3033 | -106.0265
16.3069 | -106.0483
16.3106 | -106.0688
16.3143 | -106.0877
16.3179 | -106.1093
16.3216 | -106.1338
16.3253 | -106.1574
16.3289 | -106.181
16.3326 | -106.2003
16.3362 | -106.2205
16.3399 | -106.2366
16.3436 | -106.2517
16.3472 | -106.2662
16.3509 | -106.2795
16.3546 | -106.293
16.3582 | -106.3064
16.3619 | -106.3178
16.3656 | -106.3276
16.3692 | -106.3333
16.3729 | -106.3345
16.3765 | -106.3349
16.3802 | -106.3276
16.3839 | -106.327
16.3875 | -106.3256
16.3912 | -106.3251
16.3949 | -106.3217
16.3985 | -106.3171
16.4022 | -106.3137
16.4058 | -106.3096
16.4095 | -106.3063
16.4132 | -106.3007
16.4168 | -106.2966
16.4205 | -106.2954
16.4242 | -106.2921
16.4278 | -106.291
16.4315 | -106.291
16.4351 | -106.2881
16.4388 | -106.2878
16.4425 | -106.2872
16.4461 | -106.2873
16.4498 | -106.2922
16.4535 | -106.3001
16.4571 | -106.3088
16.4608 | -106.3152
16.4644 | -106.3229
16.4681 | -106.3325
16.4718 | -106.3441
16.4754 | -106.3548
16.4791 | -106.3655
16.4828 | -106.3747
16.4864 | -106.385
16.4901 | -106.3965
16.4937 | -106.4042
16.4974 | -106.4115
16.5011 | -106.4214
16.5047 | -106.4298
16.5084 | -106.4395
16.5121 | -106.4449
16.5157 | -106.4469
16.5194 | -106.4519
16.5231 | -106.4564
16.5267 | -106.4582
16.5304 | -106.4585
16.534 | -106.457
16.5377 | -106.4537
16.5414 | -106.4486
16.545 | -106.4452
16.5487 | -106.4408
16.5524 | -106.434
16.556 | -106.4247
16.5597 | -106.4135
16.5633 | -106.4021
16.567 | -106.3946
16.5707 | -106.385
16.5743 | -106.3764
16.578 | -106.3665
16.5817 | -106.357
16.5853 | -106.349
16.589 | -106.3418
16.5926 | -106.335
16.5963 | -106.3305
16.6 | -106.3282
16.6036 | -106.3289
16.6073 | -106.3318
16.611 | -106.3355
16.6146 | -106.337
16.6183 | -106.3382
16.6219 | -106.3431
16.6256 | -106.3486
16.6293 | -106.3562
16.6329 | -106.3652
16.6366 | -106.3739
16.6403 | -106.3841
16.6439 | -106.3964
16.6476 | -106.4085
16.6513 | -106.4228
16.6549 | -106.4355
16.6586 | -106.4462
16.6622 | -106.4572
16.6659 | -106.4661
16.6696 | -106.4747
16.6732 | -106.4848
16.6769 | -106.4919
16.6806 | -106.4993
16.6842 | -106.5058
16.6879 | -106.5057
16.6915 | -106.5036
16.6952 | -106.5037
16.6989 | -106.5054
16.7025 | -106.5045
16.7062 | -106.5018
16.7099 | -106.4987
16.7135 | -106.495
16.7172 | -106.492
16.7208 | -106.4883
16.7245 | -106.4811
16.7282 | -106.4734
16.7318 | -106.4707
16.7355 | -106.4691
16.7392 | -106.4623
16.7428 | -106.4571
16.7465 | -106.4513
16.7501 | -106.4451
16.7538 | -106.4399
16.7575 | -106.4343
16.7611 | -106.4313
16.7648 | -106.4291
16.7685 | -106.4246
16.7721 | -106.4222
16.7758 | -106.4223
16.7795 | -106.4217
16.7831 | -106.4205
16.7868 | -106.4197
16.7904 | -106.4211
16.7941 | -106.4228
16.7978 | -106.4268
16.8014 | -106.4305
16.8051 | -106.4317
16.8088 | -106.4348
16.8124 | -106.4405
16.8161 | -106.4467
16.8197 | -106.451
16.8234 | -106.4551
16.8271 | -106.4586
16.8307 | -106.4609
16.8344 | -106.4618
16.8381 | -106.4638
16.8417 | -106.466
16.8454 | -106.4675
16.849 | -106.4691
16.8527 | -106.473
16.8564 | -106.4771
16.86 | -106.4778
16.8637 | -106.4799
16.8674 | -106.4803
16.871 | -106.4808
16.8747 | -106.482
16.8783 | -106.4854
16.882 | -106.4863
16.8857 | -106.4858
16.8893 | -106.4847
16.893 | -106.4854
16.8967 | -106.4884
16.9003 | -106.4908
16.904 | -106.4917
16.9077 | -106.4915
16.9113 | -106.4931
16.915 | -106.4937
16.9186 | -106.4963
16.9223 | -106.4961
16.926 | -106.4959
16.9296 | -106.4971
16.9333 | -106.4948
16.937 | -106.4912
16.9406 | -106.4871
16.9443 | -106.4849
16.9479 | -106.4829
16.9516 | -106.4799
16.9553 | -106.4769
16.9589 | -106.4729
16.9626 | -106.4687
16.9663 | -106.4655
16.9699 | -106.4617
16.9736 | -106.4543
16.9772 | -106.4472
16.9809 | -106.4419
16.9846 | -106.4362
16.9882 | -106.4322
16.9919 | -106.43
16.9956 | -106.4263
16.9992 | -106.4235
17.0029 | -106.4183
17.0065 | -106.4134
17.0102 | -106.4118
17.0139 | -106.4114
17.0175 | -106.4084
17.0212 | -106.4087
17.0249 | -106.4104
17.0285 | -106.4127
17.0322 | -106.4151
17.0358 | -106.4187
17.0395 | -106.4205
17.0432 | -106.421
17.0468 | -106.4224
17.0505 | -106.4261
17.0542 | -106.4311
17.0578 | -106.4344
17.0615 | -106.4406
17.0652 | -106.4481
17.0688 | -106.4564
17.0725 | -106.4668
17.0761 | -106.4745
17.0798 | -106.4822
17.0835 | -106.4875
17.0871 | -106.4947
17.0908 | -106.5012
17.0945 | -106.5057
17.0981 | -106.5084
17.1018 | -106.5112
17.1054 | -106.5117
17.1091 | -106.5122
17.1128 | -106.5106
17.1164 | -106.5081
17.1201 | -106.5042
17.1238 | -106.4997
17.1274 | -106.4965
17.1311 | -106.4933
17.1347 | -106.4879
17.1384 | -106.4807
17.1421 | -106.4737
17.1457 | -106.467
17.1494 | -106.4605
17.1531 | -106.4519
17.1567 | -106.4416
17.1604 | -106.4317
17.164 | -106.4223
17.1677 | -106.4134
17.1714 | -106.4042
17.175 | -106.4005
17.1787 | -106.3984
17.1824 | -106.3944
17.186 | -106.3891
17.1897 | -106.3827
17.1934 | -106.3787
17.197 | -106.3744
17.2007 | -106.3692
17.2043 | -106.3654
17.208 | -106.3622
17.2117 | -106.361
17.2153 | -106.3609
17.219 | -106.3608
17.2227 | -106.3575
17.2263 | -106.3549
17.23 | -106.352
17.2336 | -106.3488
17.2373 | -106.346
17.241 | -106.3441
17.2446 | -106.339
17.2483 | -106.3339
17.252 | -106.3309
17.2556 | -106.3278
17.2593 | -106.3189
17.2629 | -106.3089
17.2666 | -106.299
17.2703 | -106.2895
17.2739 | -106.2792
17.2776 | -106.2683
17.2813 | -106.2566
17.2849 | -106.2459
17.2886 | -106.2351
17.2922 | -106.222
17.2959 | -106.2079
17.2996 | -106.1905
17.3032 | -106.173
17.3069 | -106.1542
17.3106 | -106.1367
17.3142 | -106.1221
17.3179 | -106.1062
17.3216 | -106.092
17.3252 | -106.0768
17.3289 | -106.0627
17.3325 | -106.0493
17.3362 | -106.0378
17.3399 | -106.0272
17.3435 | -106.0207
17.3472 | -106.0127
17.3509 | -106.0055
17.3545 | -105.9996
17.3582 | -105.9951
17.3618 | -105.992
17.3655 | -105.9912
17.3692 | -105.9913
17.3728 | -105.9935
17.3765 | -105.9964
17.3802 | -106.0002
17.3838 | -106.0045
17.3875 | -106.0112
17.3911 | -106.0193
17.3948 | -106.0282
17.3985 | -106.0392
17.4021 | -106.051
17.4058 | -106.0646
17.4095 | -106.0774
17.4131 | -106.0892
17.4168 | -106.1022
17.4204 | -106.1135
17.4241 | -106.1274
17.4278 | -106.1427
17.4314 | -106.1579
17.4351 | -106.1711
17.4388 | -106.1847
17.4424 | -106.1976
17.4461 | -106.2112
17.4498 | -106.2253
17.4534 | -106.2366
17.4571 | -106.2458
17.4607 | -106.2517
17.4644 | -106.2582
17.4681 | -106.2649
17.4717 | -106.2711
17.4754 | -106.2766
17.4791 | -106.2809
17.4827 | -106.2841
17.4864 | -106.2859
17.49 | -106.2853
17.4937 | -106.2844
17.4974 | -106.2821
17.501 | -106.2785
17.5047 | -106.2713
17.5084 | -106.2673
17.512 | -106.2638
17.5157 | -106.2556
17.5193 | -106.2434
17.523 | -106.2314
17.5267 | -106.217
17.5303 | -106.2008
17.534 | -106.1837
17.5377 | -106.1667
17.5413 | -106.1515
17.545 | -106.1319
17.5486 | -106.1139
17.5523 | -106.0981
17.556 | -106.0795
17.5596 | -106.0566
17.5633 | -106.0341
17.567 | -106.0135
17.5706 | -105.996
17.5743 | -105.9788
17.578 | -105.9627
17.5816 | -105.947
17.5853 | -105.9299
17.5889 | -105.9136
17.5926 | -105.8974
17.5963 | -105.8855
17.5999 | -105.8729
17.6036 | -105.8618
17.6073 | -105.8528
17.6109 | -105.8428
17.6146 | -105.8335
17.6182 | -105.8293
17.6219 | -105.8246
17.6256 | -105.8207
17.6292 | -105.819
17.6329 | -105.8178
17.6366 | -105.8195
17.6402 | -105.8227
17.6439 | -105.8259
17.6475 | -105.8317
17.6512 | -105.8362
17.6549 | -105.8434
17.6585 | -105.8507
17.6622 | -105.8592
17.6659 | -105.8678
17.6695 | -105.8769
17.6732 | -105.8867
17.6768 | -105.8969
17.6805 | -105.9064
17.6842 | -105.9157
17.6878 | -105.929
17.6915 | -105.9426
17.6952 | -105.9553
17.6988 | -105.9675
17.7025 | -105.9792
17.7061 | -105.9918
17.7098 | -106.0038
17.7135 | -106.0154
17.7171 | -106.0251
17.7208 | -106.0331
17.7245 | -106.0424
17.7281 | -106.0522
17.7318 | -106.0625
17.7355 | -106.0703
17.7391 | -106.0759
17.7428 | -106.0854
17.7464 | -106.0938
17.7501 | -106.1015
17.7538 | -106.1075
17.7574 | -106.1151
17.7611 | -106.1203
17.7648 | -106.1272
17.7684 | -106.1343
17.7721 | -106.141
17.7757 | -106.1464
17.7794 | -106.152
17.7831 | -106.1599
17.7867 | -106.1661
17.7904 | -106.171
17.7941 | -106.1728
17.7977 | -106.1784
17.8014 | -106.1843
17.805 | -106.1889
17.8087 | -106.1941
17.8124 | -106.2013
17.816 | -106.2053
17.8197 | -106.2081
17.8234 | -106.2091
17.827 | -106.2117
17.8307 | -106.2163
17.8343 | -106.2192
17.838 | -106.223
17.8417 | -106.2276
17.8453 | -106.231
17.849 | -106.2342
17.8527 | -106.2395
17.8563 | -106.2456
17.86 | -106.2516
17.8637 | -106.2602
17.8673 | -106.2673
17.871 | -106.2735
17.8746 | -106.2812
17.8783 | -106.2894
17.882 | -106.2986
17.8856 | -106.3083
17.8893 | -106.3143
17.893 | -106.3216
17.8966 | -106.3289
17.9003 | -106.3368
17.9039 | -106.3476
17.9076 | -106.3562
17.9113 | -106.3657
17.9149 | -106.3753
17.9186 | -106.3843
17.9223 | -106.3913
17.9259 | -106.3962
17.9296 | -106.402
17.9332 | -106.4089
17.9369 | -106.4126
17.9406 | -106.4154
17.9442 | -106.4166
17.9479 | -106.4172
17.9516 | -106.4189
17.9552 | -106.4232
17.9589 | -106.4289
17.9625 | -106.4339
17.9662 | -106.4336
17.9699 | -106.4373
17.9735 | -106.4404
17.9772 | -106.4445
17.9809 | -106.4479
17.9845 | -106.4499
17.9882 | -106.4508
17.9918 | -106.4508
17.9955 | -106.4537
17.9992 | -106.4547
18.0028 | -106.4559
18.0065 | -106.4572
18.0102 | -106.4558
18.0138 | -106.4575
18.0175 | -106.4594
18.0212 | -106.4591
18.0248 | -106.4604
18.0285 | -106.4618
18.0321 | -106.4617
18.0358 | -106.4588
18.0395 | -106.4544
18.0431 | -106.4514
18.0468 | -106.4476
18.0505 | -106.4439
18.0541 | -106.4375
18.0578 | -106.4284
18.0614 | -106.4201
18.0651 | -106.411
18.0688 | -106.3984
18.0724 | -106.3863
18.0761 | -106.3751
18.0798 | -106.3646
18.0834 | -106.354
18.0871 | -106.3396
18.0907 | -106.3272
18.0944 | -106.3159
18.0981 | -106.3024
18.1017 | -106.2893
18.1054 | -106.2767
18.1091 | -106.264
18.1127 | -106.2487
18.1164 | -106.2349
18.12 | -106.2211
18.1237 | -106.2101
18.1274 | -106.1976
18.131 | -106.187
18.1347 | -106.1782
18.1384 | -106.1698
18.142 | -106.1629
18.1457 | -106.1567
18.1494 | -106.1503
18.153 | -106.1461
18.1567 | -106.1408
18.1603 | -106.1364
18.164 | -106.1353
18.1677 | -106.1351
18.1713 | -106.134
18.175 | -106.1329
18.1787 | -106.1334
18.1823 | -106.136
18.186 | -106.138
18.1896 | -106.1414
18.1933 | -106.1489
18.197 | -106.1542
18.2006 | -106.1596
18.2043 | -106.166
18.208 | -106.171
18.2116 | -106.1742
18.2153 | -106.1756
18.2189 | -106.1799
18.2226 | -106.1838
18.2263 | -106.19
18.2299 | -106.1954
18.2336 | -106.2001
18.2373 | -106.2038
18.2409 | -106.2071
18.2446 | -106.2096
18.2482 | -106.2145
18.2519 | -106.2205
18.2556 | -106.224
18.2592 | -106.2264
18.2629 | -106.2293
18.2666 | -106.2321
18.2702 | -106.235
18.2739 | -106.2345
18.2776 | -106.2331
18.2812 | -106.2326
18.2849 | -106.2328
18.2885 | -106.2324
18.2922 | -106.2328
18.2959 | -106.234
18.2995 | -106.2351
18.3032 | -106.236
18.3069 | -106.2392
18.3105 | -106.2423
18.3142 | -106.2429
18.3178 | -106.247
18.3215 | -106.2507
18.3252 | -106.2555
18.3288 | -106.2595
18.3325 | -106.2598
18.3362 | -106.2614
18.3398 | -106.2642
18.3435 | -106.266
18.3471 | -106.2687
18.3508 | -106.2692
18.3545 | -106.2695
18.3581 | -106.2734
18.3618 | -106.2731
18.3655 | -106.275
18.3691 | -106.2729
18.3728 | -106.2718
18.3764 | -106.2727
18.3801 | -106.2728
18.3838 | -106.2715
18.3874 | -106.2709
18.3911 | -106.2684
18.3948 | -106.2633
18.3984 | -106.2588
18.4021 | -106.2527
18.4058 | -106.2471
18.4094 | -106.2403
18.4131 | -106.2307
18.4167 | -106.222
18.4204 | -106.2131
18.4241 | -106.2003
18.4277 | -106.1893
18.4314 | -106.18
18.4351 | -106.1702
18.4387 | -106.1604
18.4424 | -106.148
18.446 | -106.1371
18.4497 | -106.1275
18.4534 | -106.1153
18.457 | -106.1057
18.4607 | -106.098
18.4644 | -106.0905
18.468 | -106.0844
18.4717 | -106.0786
18.4753 | -106.0751
18.479 | -106.0704
18.4827 | -106.0692
18.4863 | -106.0688
18.49 | -106.0701
18.4937 | -106.0825
18.4973 | -106.084
18.501 | -106.0875
18.5046 | -106.0911
18.5083 | -106.0963
18.512 | -106.1025
18.5156 | -106.1111
18.5193 | -106.1215
18.523 | -106.1329
18.5266 | -106.1435
18.5303 | -106.1535
18.534 | -106.1656
18.5376 | -106.1793
18.5413 | -106.1937
18.5449 | -106.2071
18.5486 | -106.2204
18.5523 | -106.2353
18.5559 | -106.2481
18.5596 | -106.2607
18.5633 | -106.2719
18.5669 | -106.2823
18.5706 | -106.292
18.5742 | -106.3024
18.5779 | -106.3152
18.5816 | -106.3266
18.5852 | -106.3351
18.5889 | -106.3435
18.5926 | -106.3524
18.5962 | -106.3611
18.5999 | -106.3664
18.6035 | -106.3706
18.6072 | -106.3747
18.6109 | -106.3793
18.6145 | -106.3829
18.6182 | -106.3841
18.6219 | -106.3819
18.6255 | -106.3792
18.6292 | -106.3731
18.6328 | -106.3688
18.6365 | -106.364
18.6402 | -106.3548
18.6438 | -106.3453
18.6475 | -106.3348
18.6512 | -106.3224
18.6548 | -106.3098
18.6585 | -106.2988
18.6621 | -106.2867
18.6658 | -106.2751
18.6695 | -106.2627
18.6731 | -106.2512
18.6768 | -106.2405
18.6805 | -106.2287
18.6841 | -106.2178
18.6878 | -106.2067
18.6915 | -106.1949
18.6951 | -106.1842
18.6988 | -106.174
18.7024 | -106.1651
18.7061 | -106.159
18.7098 | -106.1533
18.7134 | -106.1469
18.7171 | -106.1443
18.7208 | -106.142
18.7244 | -106.1397
18.7281 | -106.138
18.7317 | -106.1377
18.7354 | -106.1401
18.7391 | -106.1428
18.7427 | -106.1468
18.7464 | -106.1502
18.7501 | -106.1533
18.7537 | -106.156
18.7574 | -106.1592
18.761 | -106.1651
18.7647 | -106.1717
18.7684 | -106.1776
18.772 | -106.1858
18.7757 | -106.1931
18.7794 | -106.2004
18.783 | -106.2068
18.7867 | -106.2131
18.7903 | -106.2201
18.794 | -106.23
18.7977 | -106.239
18.8013 | -106.2472
18.805 | -106.2535
18.8087 | -106.259
18.8123 | -106.2634
18.816 | -106.2679
18.8197 | -106.2717
18.8233 | -106.2769
18.827 | -106.2811
18.8306 | -106.2861
18.8343 | -106.289
18.838 | -106.2915
18.8416 | -106.2946
18.8453 | -106.2964
18.849 | -106.2962
18.8526 | -106.2972
18.8563 | -106.2962
18.8599 | -106.2969
18.8636 | -106.2987
18.8673 | -106.2989
18.8709 | -106.296
18.8746 | -106.2939
18.8783 | -106.2918
18.8819 | -106.2923
18.8856 | -106.2917
18.8892 | -106.2892
18.8929 | -106.2913
18.8966 | -106.2927
18.9002 | -106.2952
18.9039 | -106.2981
18.9076 | -106.3015
18.9112 | -106.3059
18.9149 | -106.3098
18.9185 | -106.3161
18.9222 | -106.3207
18.9259 | -106.3244
18.9295 | -106.3291
18.9332 | -106.3345
18.9369 | -106.339
18.9405 | -106.345
18.9442 | -106.3527
18.9478 | -106.3601
18.9515 | -106.3669
18.9552 | -106.3749
18.9588 | -106.3795
18.9625 | -106.3835
18.9662 | -106.3871
18.9698 | -106.3888
18.9735 | -106.3913
18.9772 | -106.3935
18.9808 | -106.3957
18.9845 | -106.3973
18.9881 | -106.3986
18.9918 | -106.4001
18.9955 | -106.4018
18.9991 | -106.4018
19.0028 | -106.3999
19.0065 | -106.3974
19.0101 | -106.3931
19.0138 | -106.3893
19.0174 | -106.3876
19.0211 | -106.3772
19.0248 | -106.37
19.0284 | -106.3649
19.0321 | -106.3595
19.0358 | -106.3527
19.0394 | -106.3469
19.0431 | -106.3381
19.0467 | -106.3314
19.0504 | -106.3239
19.0541 | -106.3156
19.0577 | -106.3075
19.0614 | -106.3005
19.0651 | -106.2917
19.0687 | -106.2851
19.0724 | -106.279
19.076 | -106.2744
19.0797 | -106.2712
19.0834 | -106.2673
19.087 | -106.2591
19.0907 | -106.2527
19.0944 | -106.2492
19.098 | -106.2474
19.1017 | -106.2461
19.1054 | -106.2469
19.109 | -106.2467
19.1127 | -106.2452
19.1163 | -106.2461
19.12 | -106.2485
19.1237 | -106.252
19.1273 | -106.2568
19.131 | -106.2609
19.1347 | -106.263
19.1383 | -106.267
19.142 | -106.2736
19.1456 | -106.2801
19.1493 | -106.2875
19.153 | -106.2934
19.1566 | -106.2997
19.1603 | -106.3091
19.164 | -106.317
19.1676 | -106.3233
19.1713 | -106.3281
19.1749 | -106.3338
19.1786 | -106.3395
19.1823 | -106.3427
19.1859 | -106.347
19.1896 | -106.3527
19.1933 | -106.3599
19.1969 | -106.3623
19.2006 | -106.3648
19.2042 | -106.3672
19.2079 | -106.3693
19.2116 | -106.3692
19.2152 | -106.369
19.2189 | -106.3678
19.2226 | -106.3658
19.2262 | -106.362
19.2299 | -106.3604
19.2336 | -106.3573
19.2372 | -106.3514
19.2409 | -106.344
19.2445 | -106.3348
19.2482 | -106.3239
19.2519 | -106.3118
19.2555 | -106.2997
19.2592 | -106.2883
19.2629 | -106.2773
19.2665 | -106.2675
19.2702 | -106.2577
19.2738 | -106.2478
19.2775 | -106.2387
19.2812 | -106.2309
19.2848 | -106.2239
19.2885 | -106.2145
19.2922 | -106.206
19.2958 | -106.2004
19.2995 | -106.1971
19.3031 | -106.1913
19.3068 | -106.1867
19.3105 | -106.1828
19.3141 | -106.1797
19.3178 | -106.1794
19.3215 | -106.1789
19.3251 | -106.1774
19.3288 | -106.1789
19.3324 | -106.1832
19.3361 | -106.1879
19.3398 | -106.1937
19.3434 | -106.2006
19.3471 | -106.2075
19.3508 | -106.2165
19.3544 | -106.2265
19.3581 | -106.239
19.3618 | -106.2544
19.3654 | -106.2693
19.3691 | -106.2821
19.3727 | -106.2938
19.3764 | -106.3052
19.3801 | -106.3176
19.3837 | -106.3324
19.3874 | -106.3478
19.3911 | -106.3615
19.3947 | -106.3744
19.3984 | -106.3889
19.402 | -106.4036
19.4057 | -106.4173
19.4094 | -106.43
19.413 | -106.4418
19.4167 | -106.4515
19.4204 | -106.4604
19.424 | -106.4669
19.4277 | -106.4745
19.4313 | -106.4827
19.435 | -106.4846
19.4387 | -106.4842
19.4423 | -106.4861
19.446 | -106.4867
19.4497 | -106.4838
19.4533 | -106.4796
19.457 | -106.4779
19.4606 | -106.4732
19.4643 | -106.4651
19.468 | -106.4561
19.4716 | -106.4495
19.4753 | -106.4378
19.479 | -106.4253
19.4826 | -106.4133
19.4863 | -106.4009
19.49 | -106.3888
19.4936 | -106.3769
19.4973 | -106.3662
19.5009 | -106.3567
19.5046 | -106.3472
19.5083 | -106.3368
19.5119 | -106.3272
19.5156 | -106.3191
19.5193 | -106.309
19.5229 | -106.2992
19.5266 | -106.2905
19.5302 | -106.28
19.5339 | -106.2701
19.5376 | -106.2606
19.5412 | -106.2502
19.5449 | -106.2421
19.5486 | -106.2417
19.5522 | -106.2356
19.5559 | -106.2305
19.5595 | -106.2279
19.5632 | -106.2247
19.5669 | -106.2213
19.5705 | -106.2207
19.5742 | -106.218
19.5779 | -106.2129
19.5815 | -106.2101
19.5852 | -106.2073
19.5888 | -106.2073
19.5925 | -106.2052
19.5962 | -106.2018
19.5998 | -106.1982
19.6035 | -106.1952
19.6072 | -106.1923
19.6108 | -106.188
19.6145 | -106.1835
19.6181 | -106.1786
19.6218 | -106.1756
19.6255 | -106.1729
19.6291 | -106.1682
19.6328 | -106.1621
19.6365 | -106.1571
19.6401 | -106.1537
19.6438 | -106.1497
19.6475 | -106.1452
19.6511 | -106.1413
19.6548 | -106.138
19.6584 | -106.133
19.6621 | -106.1278
19.6658 | -106.1228
19.6694 | -106.1175
19.6731 | -106.1141
19.6768 | -106.108
19.6804 | -106.1015
19.6841 | -106.0962
19.6877 | -106.0953
19.6914 | -106.0912
19.6951 | -106.0876
19.6987 | -106.0846
19.7024 | -106.0827
19.7061 | -106.0783
19.7097 | -106.0756
19.7134 | -106.0759
19.717 | -106.0757
19.7207 | -106.0751
19.7244 | -106.0764
19.728 | -106.0773
19.7317 | -106.078
19.7354 | -106.0777
19.739 | -106.0783
19.7427 | -106.0793
19.7463 | -106.0803
19.75 | -106.0848
19.7537 | -106.0893
19.7573 | -106.0912
19.761 | -106.0936
19.7647 | -106.096
19.7683 | -106.0997
19.772 | -106.1042
19.7757 | -106.1077
19.7793 | -106.1115
19.783 | -106.1148
19.7866 | -106.1202
19.7903 | -106.125
19.794 | -106.1295
19.7976 | -106.1343
19.8013 | -106.1387
19.805 | -106.1428
19.8086 | -106.1469
19.8123 | -106.1486
19.8159 | -106.1516
19.8196 | -106.1552
19.8233 | -106.1594
19.8269 | -106.16
19.8306 | -106.1599
19.8343 | -106.1619
19.8379 | -106.1638
19.8416 | -106.1635
19.8452 | -106.1647
19.8489 | -106.1675
19.8526 | -106.1698
19.8562 | -106.1727
19.8599 | -106.1766
19.8636 | -106.1811
19.8672 | -106.1847
19.8709 | -106.1881
19.8745 | -106.1929
19.8782 | -106.1983
19.8819 | -106.204
19.8855 | -106.2115
19.8892 | -106.2175
19.8929 | -106.2224
19.8965 | -106.2284
19.9002 | -106.2362
19.9039 | -106.2415
19.9075 | -106.2492
19.9112 | -106.2571
19.9148 | -106.2661
19.9185 | -106.2752
19.9222 | -106.2855
19.9258 | -106.2962
19.9295 | -106.3077
19.9332 | -106.3197
19.9368 | -106.3294
19.9405 | -106.3409
19.9441 | -106.3515
19.9478 | -106.3611
19.9515 | -106.3722
19.9551 | -106.3822
19.9588 | -106.3913
19.9625 | -106.401
19.9661 | -106.4106
19.9698 | -106.416
19.9734 | -106.4206
19.9771 | -106.4249
19.9808 | -106.4294
19.9844 | -106.4334
19.9881 | -106.4388
19.9918 | -106.4402
19.9954 | -106.4402
19.9991 | -106.4414
20.0027 | -106.4412
20.0064 | -106.4383
20.0101 | -106.4334
20.0137 | -106.4284
20.0174 | -106.4236
20.0211 | -106.421
20.0247 | -106.4171
20.0284 | -106.4105
20.0321 | -106.401
20.0357 | -106.3937
20.0394 | -106.3863
20.043 | -106.3767
20.0467 | -106.3681
20.0504 | -106.3566
20.054 | -106.3452
20.0577 | -106.3358
20.0614 | -106.3273
20.065 | -106.3183
20.0687 | -106.3084
20.0723 | -106.2996
20.076 | -106.2921
20.0797 | -106.29
20.0833 | -106.2834
20.087 | -106.2776
20.0907 | -106.2725
20.0943 | -106.2681
20.098 | -106.2653
20.1016 | -106.2614
20.1053 | -106.2576
20.109 | -106.2539
20.1126 | -106.2503
20.1163 | -106.2484
20.12 | -106.248
20.1236 | -106.247
20.1273 | -106.2467
20.1309 | -106.2453
20.1346 | -106.2441
20.1383 | -106.244
20.1419 | -106.2475
20.1456 | -106.2517
20.1493 | -106.254
20.1529 | -106.2562
20.1566 | -106.2583
20.1602 | -106.2604
20.1639 | -106.2631
20.1676 | -106.2668
20.1712 | -106.269
20.1749 | -106.2705
20.1786 | -106.2749
20.1822 | -106.2776
20.1859 | -106.2801
20.1896 | -106.2812
20.1932 | -106.2807
20.1969 | -106.2785
20.2005 | -106.2779
20.2042 | -106.2771
20.2079 | -106.2755
20.2115 | -106.2725
20.2152 | -106.2705
20.2189 | -106.2681
20.2225 | -106.2668
20.2262 | -106.2641
20.2298 | -106.2616
20.2335 | -106.2613
20.2372 | -106.2589
20.2408 | -106.256
20.2445 | -106.2547
20.2482 | -106.2521
20.2518 | -106.249
20.2555 | -106.2451
20.2591 | -106.2419
20.2628 | -106.2378
20.2665 | -106.2327
20.2701 | -106.2266
20.2738 | -106.2183
20.2775 | -106.2108
20.2811 | -106.203
20.2848 | -106.1945
20.2884 | -106.1858
20.2921 | -106.179
20.2958 | -106.1746
20.2994 | -106.1683
20.3031 | -106.1606
20.3068 | -106.1508
20.3104 | -106.1405
20.3141 | -106.1301
20.3178 | -106.1187
20.3214 | -106.1081
20.3251 | -106.0996
20.3287 | -106.0918
20.3324 | -106.0817
20.3361 | -106.0701
20.3397 | -106.0611
20.3434 | -106.0517
20.3471 | -106.0423
20.3507 | -106.0338
20.3544 | -106.0254
20.358 | -106.0168
20.3617 | -106.0076
20.3654 | -105.999
20.369 | -105.993
20.3727 | -105.9861
20.3764 | -105.9798
20.38 | -105.9733
20.3837 | -105.9672
20.3873 | -105.9629
20.391 | -105.9589
20.3947 | -105.955
20.3983 | -105.9506
20.402 | -105.9488
20.4057 | -105.9471
20.4093 | -105.9434
20.413 | -105.9425
20.4166 | -105.9427
20.4203 | -105.9442
20.424 | -105.9457
20.4276 | -105.9464
20.4313 | -105.9454
20.435 | -105.9481
20.4386 | -105.951
20.4423 | -105.9541
20.4459 | -105.9577
20.4496 | -105.9627
20.4533 | -105.9665
20.4569 | -105.9685
20.4606 | -105.9721
20.4643 | -105.9781
20.4679 | -105.9826
20.4716 | -105.9881
20.4753 | -105.9941
20.4789 | -105.9994
20.4826 | -106.0035
20.4862 | -106.0112
20.4899 | -106.0186
20.4936 | -106.0229
20.4972 | -106.0268
20.5009 | -106.031
20.5046 | -106.0356
20.5082 | -106.0409
20.5119 | -106.0448
20.5155 | -106.0482
20.5192 | -106.0508
20.5229 | -106.0531
20.5265 | -106.0549
20.5302 | -106.0558
20.5339 | -106.0563
20.5375 | -106.0563
20.5412 | -106.0554
20.5448 | -106.0562
20.5485 | -106.0562
20.5522 | -106.057
20.5558 | -106.0563
20.5595 | -106.0549
20.5632 | -106.0531
20.5668 | -106.0511
20.5705 | -106.0498
20.5741 | -106.0474
20.5778 | -106.0434
20.5815 | -106.0383
20.5851 | -106.0318
20.5888 | -106.0273
20.5925 | -106.0216
20.5961 | -106.014
20.5998 | -106.0091
20.6035 | -106.0068
20.6071 | -106.0015
20.6108 | -105.9963
20.6144 | -105.9885
20.6181 | -105.9807
20.6218 | -105.9712
20.6254 | -105.965
20.6291 | -105.9603
20.6328 | -105.9518
20.6364 | -105.9434
20.6401 | -105.9344
20.6437 | -105.9269
20.6474 | -105.9186
20.6511 | -105.9112
20.6547 | -105.9033
20.6584 | -105.894
20.6621 | -105.8838
20.6657 | -105.8758
20.6694 | -105.8673
20.673 | -105.8583
20.6767 | -105.8511
20.6804 | -105.8413
20.684 | -105.8325
20.6877 | -105.8235
20.6914 | -105.8152
20.695 | -105.8072
20.6987 | -105.7986
20.7023 | -105.7919
20.706 | -105.7837
20.7097 | -105.7763
20.7133 | -105.7707
20.717 | -105.766
20.7207 | -105.7621
20.7243 | -105.7581
20.728 | -105.7515
20.7317 | -105.7479
20.7353 | -105.7468
20.739 | -105.746
20.7426 | -105.7441
20.7463 | -105.7452
20.75 | -105.7475
20.7536 | -105.7494
20.7573 | -105.7507
20.761 | -105.7533
20.7646 | -105.7549
20.7683 | -105.7568
20.7719 | -105.7632
20.7756 | -105.7708
20.7793 | -105.7761
20.7829 | -105.7796
20.7866 | -105.7854
20.7903 | -105.7937
20.7939 | -105.8046
20.7976 | -105.8114
20.8012 | -105.8193
20.8049 | -105.8285
20.8086 | -105.8387
20.8122 | -105.8503
20.8159 | -105.8597
20.8196 | -105.8708
20.8232 | -105.8831
20.8269 | -105.8955
20.8305 | -105.9073
20.8342 | -105.9188
20.8379 | -105.9298
20.8415 | -105.9407
20.8452 | -105.949
20.8489 | -105.9575
20.8525 | -105.9681
20.8562 | -105.9796
20.8599 | -105.9888
20.8635 | -105.9961
20.8672 | -106.0017
20.8708 | -106.0078
20.8745 | -106.0136
20.8782 | -106.0208
20.8818 | -106.0249
20.8855 | -106.0291
20.8892 | -106.0347
20.8928 | -106.0367
20.8965 | -106.0369
20.9001 | -106.0352
20.9038 | -106.035
20.9075 | -106.0364
20.9111 | -106.0364
20.9148 | -106.0331
20.9185 | -106.0324
20.9221 | -106.0283
20.9258 | -106.0224
20.9294 | -106.0195
20.9331 | -106.0151
20.9368 | -106.0119
20.9404 | -106.0085
20.9441 | -106.0044
20.9478 | -105.9997
20.9514 | -105.9972
20.9551 | -105.9951
20.9587 | -105.9921
20.9624 | -105.9917
20.9661 | -105.9923
20.9697 | -105.9929
20.9734 | -105.993
20.9771 | -105.9959
20.9807 | -105.9983
20.9844 | -106.0017
20.9881 | -106.0057
20.9917 | -106.0087
20.9954 | -106.0122
20.999 | -106.0139
21.0027 | -106.0168
21.0064 | -106.0208
21.01 | -106.0244
21.0137 | -106.0282
21.0174 | -106.0336
21.021 | -106.0386
21.0247 | -106.046
21.0283 | -106.0515
21.032 | -106.0559
21.0357 | -106.0607
21.0393 | -106.0674
21.043 | -106.0728
21.0467 | -106.0777
21.0503 | -106.0833
21.054 | -106.0896
21.0576 | -106.0963
21.0613 | -106.1031
21.065 | -106.1095
21.0686 | -106.114
21.0723 | -106.1176
21.076 | -106.1209
21.0796 | -106.1246
21.0833 | -106.1294
21.0869 | -106.1323
21.0906 | -106.1331
21.0943 | -106.1351
21.0979 | -106.1343
21.1016 | -106.1357
21.1053 | -106.1358
21.1089 | -106.1336
21.1126 | -106.1323
21.1162 | -106.1303
21.1199 | -106.1283
21.1236 | -106.1245
21.1272 | -106.1226
21.1309 | -106.1208
21.1346 | -106.1126
21.1382 | -106.1087
21.1419 | -106.1046
21.1456 | -106.1001
21.1492 | -106.0938
21.1529 | -106.0864
21.1565 | -106.0777
21.1602 | -106.0704
21.1639 | -106.0625
21.1675 | -106.0557
21.1712 | -106.0489
21.1749 | -106.0415
21.1785 | -106.0331
21.1822 | -106.0276
21.1858 | -106.0214
21.1895 | -106.014
21.1932 | -106.0061
21.1968 | -105.9978
21.2005 | -105.9911
21.2042 | -105.9837
21.2078 | -105.9769
21.2115 | -105.9716
21.2151 | -105.9679
21.2188 | -105.9624
21.2225 | -105.9569
21.2261 | -105.9514
21.2298 | -105.9463
21.2335 | -105.9428
21.2371 | -105.9388
21.2408 | -105.9354
21.2444 | -105.9322
21.2481 | -105.9287
21.2518 | -105.9249
21.2554 | -105.9199
21.2591 | -105.9154
21.2628 | -105.9115
21.2664 | -105.9092
21.2701 | -105.9077
21.2738 | -105.9057
21.2774 | -105.9041
21.2811 | -105.9031
21.2847 | -105.9038
21.2884 | -105.9035
21.2921 | -105.9031
21.2957 | -105.9041
21.2994 | -105.9061
21.3031 | -105.9074
21.3067 | -105.9074
21.3104 | -105.9055
21.314 | -105.9065
21.3177 | -105.9088
21.3214 | -105.9115
21.325 | -105.9133
21.3287 | -105.9155
21.3324 | -105.9174
21.336 | -105.9187
21.3397 | -105.9187
21.3433 | -105.9201
21.347 | -105.9241
21.3507 | -105.9282
21.3543 | -105.9293
21.358 | -105.9291
21.3617 | -105.9317
21.3653 | -105.9346
21.369 | -105.9373
21.3726 | -105.9411
21.3763 | -105.9463
21.38 | -105.9512
21.3836 | -105.9547
21.3873 | -105.9574
21.391 | -105.9594
21.3946 | -105.9634
21.3983 | -105.9685
21.4019 | -105.9723
21.4056 | -105.976
21.4093 | -105.9802
21.4129 | -105.983
21.4166 | -105.9839
21.4203 | -105.9836
21.4239 | -105.9853
21.4276 | -105.9829
21.4313 | -105.9783
21.4349 | -105.9799
21.4386 | -105.9811
21.4422 | -105.9781
21.4459 | -105.9761
21.4496 | -105.9727
21.4532 | -105.9673
21.4569 | -105.962
21.4606 | -105.9572
21.4642 | -105.951
21.4679 | -105.9464
21.4715 | -105.9395
21.4752 | -105.9292
21.4789 | -105.9174
21.4825 | -105.9064
21.4862 | -105.8929
21.4899 | -105.8797
21.4935 | -105.868
21.4972 | -105.8567
21.5008 | -105.8453
21.5045 | -105.8323
21.5082 | -105.8185
21.5118 | -105.8033
21.5155 | -105.7899
21.5192 | -105.7764
21.5228 | -105.761
21.5265 | -105.7461
21.5301 | -105.7347
21.5338 | -105.7204
21.5375 | -105.7065
21.5411 | -105.6949
21.5448 | -105.6821
21.5485 | -105.6718
21.5521 | -105.6636
21.5558 | -105.6549
21.5595 | -105.6485
21.5631 | -105.6423
21.5668 | -105.6349
21.5704 | -105.628
21.5741 | -105.6227
21.5778 | -105.619
21.5814 | -105.6163
21.5851 | -105.6143
21.5888 | -105.6108
21.5924 | -105.6109
21.5961 | -105.6136
21.5997 | -105.616
21.6034 | -105.6202
21.6071 | -105.6267
21.6107 | -105.6333
21.6144 | -105.6401
21.6181 | -105.6453
21.6217 | -105.6486
21.6254 | -105.6524
21.629 | -105.6581
21.6327 | -105.6648
21.6364 | -105.673
21.64 | -105.6812
21.6437 | -105.6887
21.6474 | -105.6959
21.651 | -105.7029
21.6547 | -105.7087
21.6583 | -105.7142
21.662 | -105.7201
21.6657 | -105.7197
21.6693 | -105.7238
21.673 | -105.7268
21.6767 | -105.7292
21.6803 | -105.7303
21.684 | -105.7296
21.6877 | -105.7302
21.6913 | -105.73
21.695 | -105.7302
21.6986 | -105.732
21.7023 | -105.7324
21.706 | -105.7314
21.7096 | -105.7293
21.7133 | -105.7276
21.717 | -105.7249
21.7206 | -105.7237
21.7243 | -105.7226
21.7279 | -105.7224
21.7316 | -105.7195
21.7353 | -105.7197
21.7389 | -105.7211
21.7426 | -105.716
21.7463 | -105.7122
21.7499 | -105.7122
21.7536 | -105.7103
21.7572 | -105.7098
21.7609 | -105.7108
21.7646 | -105.7105
21.7682 | -105.7102
21.7719 | -105.7109
21.7756 | -105.7098
21.7792 | -105.7108
21.7829 | -105.7143
21.7865 | -105.7179
21.7902 | -105.72
21.7939 | -105.7248
21.7975 | -105.7297
21.8012 | -105.7329
21.8049 | -105.7373
21.8085 | -105.7432
21.8122 | -105.7486
21.8159 | -105.757
21.8195 | -105.7653
21.8232 | -105.7711
21.8268 | -105.7755
21.8305 | -105.7826
21.8342 | -105.7923
21.8378 | -105.7986
21.8415 | -105.807
21.8452 | -105.815
21.8488 | -105.8241
21.8525 | -105.8355
21.8561 | -105.846
21.8598 | -105.8536
21.8635 | -105.8603
21.8671 | -105.8696
21.8708 | -105.8792
21.8745 | -105.8876
21.8781 | -105.8937
21.8818 | -105.9006
21.8854 | -105.9087
21.8891 | -105.9166
21.8928 | -105.9234
21.8964 | -105.9297
21.9001 | -105.9336
21.9038 | -105.9375
21.9074 | -105.9405
21.9111 | -105.9416
21.9147 | -105.9412
21.9184 | -105.9392
21.9221 | -105.9359
21.9257 | -105.9344
21.9294 | -105.9356
21.9331 | -105.934
21.9367 | -105.9289
21.9404 | -105.9215
21.9441 | -105.913
21.9477 | -105.9046
21.9514 | -105.8977
21.955 | -105.8876
21.9587 | -105.8763
21.9624 | -105.8669
21.966 | -105.8579
21.9697 | -105.8472
21.9734 | -105.8375
21.977 | -105.8235
21.9807 | -105.809
21.9843 | -105.7964
21.988 | -105.7834
21.9917 | -105.7703
21.9953 | -105.7587
21.999 | -105.7459
22.0027 | -105.7335
22.0063 | -105.7207
22.01 | -105.708
22.0136 | -105.6963
22.0173 | -105.6848
22.021 | -105.6727
22.0246 | -105.6603
22.0283 | -105.6498
22.032 | -105.6389
22.0356 | -105.628
22.0393 | -105.6168
22.0429 | -105.6069
22.0466 | -105.5971
22.0503 | -105.5866
22.0539 | -105.58
22.0576 | -105.5698
22.0613 | -105.5606
22.0649 | -105.5535
22.0686 | -105.5463
22.0722 | -105.5394
22.0759 | -105.5331
22.0796 | -105.5264
22.0832 | -105.5202
22.0869 | -105.5156
22.0906 | -105.5109
22.0942 | -105.507
22.0979 | -105.5056
22.1016 | -105.504
22.1052 | -105.4996
22.1089 | -105.4971
22.1125 | -105.4959
22.1162 | -105.4923
22.1199 | -105.4877
22.1235 | -105.4874
22.1272 | -105.4838
22.1309 | -105.4806
22.1345 | -105.4801
22.1382 | -105.4808
22.1418 | -105.4793
22.1455 | -105.475
22.1492 | -105.4731
22.1528 | -105.4724
22.1565 | -105.4697
22.1602 | -105.4642
22.1638 | -105.459
22.1675 | -105.4549
22.1711 | -105.454
22.1748 | -105.4544
22.1785 | -105.4521
22.1821 | -105.4494
22.1858 | -105.4468
22.1895 | -105.4461
22.1931 | -105.4471
22.1968 | -105.4467
22.2004 | -105.4444
22.2041 | -105.4424
22.2078 | -105.4414
22.2114 | -105.442
22.2151 | -105.4409
22.2188 | -105.4387
22.2224 | -105.4359
22.2261 | -105.4348
22.2298 | -105.4347
22.2334 | -105.4365
22.2371 | -105.4399
22.2407 | -105.442
22.2444 | -105.4421
22.2481 | -105.4439
22.2517 | -105.447
22.2554 | -105.4493
22.2591 | -105.4522
22.2627 | -105.455
22.2664 | -105.4578
22.27 | -105.4608
22.2737 | -105.4622
22.2774 | -105.4622
22.281 | -105.4657
22.2847 | -105.4696
22.2884 | -105.474
22.292 | -105.4803
22.2957 | -105.4861
22.2993 | -105.491
22.303 | -105.4974
22.3067 | -105.503
22.3103 | -105.5072
22.314 | -105.5111
22.3177 | -105.5143
22.3213 | -105.5172
22.325 | -105.5215
22.3286 | -105.5246
22.3323 | -105.5277
22.336 | -105.5314
22.3396 | -105.5366
22.3433 | -105.542
22.347 | -105.5479
22.3506 | -105.553
22.3543 | -105.5573
22.358 | -105.5622
22.3616 | -105.5652
22.3653 | -105.5672
22.3689 | -105.568
22.3726 | -105.5705
22.3763 | -105.5749
22.3799 | -105.5786
22.3836 | -105.5797
22.3873 | -105.5804
22.3909 | -105.5826
22.3946 | -105.5844
22.3982 | -105.5845
22.4019 | -105.5853
22.4056 | -105.5858
22.4092 | -105.5838
22.4129 | -105.5808
22.4166 | -105.5799
22.4202 | -105.5813
22.4239 | -105.5815
22.4275 | -105.5803
22.4312 | -105.5812
22.4349 | -105.5821
22.4385 | -105.5802
22.4422 | -105.5791
22.4459 | -105.5767
22.4495 | -105.5722
22.4532 | -105.567
22.4568 | -105.563
22.4605 | -105.5585
22.4642 | -105.5529
22.4678 | -105.5441
22.4715 | -105.5375
22.4752 | -105.5329
22.4788 | -105.5252
22.4825 | -105.5166
22.4861 | -105.5063
22.4898 | -105.4951
22.4935 | -105.4862
22.4971 | -105.4778
22.5008 | -105.4685
22.5045 | -105.4577
22.5081 | -105.4472
22.5118 | -105.4359
22.5155 | -105.4237
22.5191 | -105.4127
22.5228 | -105.4026
22.5264 | -105.3921
22.5301 | -105.3794
22.5338 | -105.3671
22.5374 | -105.3566
22.5411 | -105.3476
22.5448 | -105.3345
22.5484 | -105.3215
22.5521 | -105.309
22.5557 | -105.2945
22.5594 | -105.2779
22.5631 | -105.2622
22.5667 | -105.2493
22.5704 | -105.2362
22.5741 | -105.2243
22.5777 | -105.2125
22.5814 | -105.2005
22.585 | -105.1867
22.5887 | -105.1721
22.5924 | -105.158
22.596 | -105.1445
22.5997 | -105.1306
22.6034 | -105.1173
22.607 | -105.1038
22.6107 | -105.0907
22.6143 | -105.0776
22.618 | -105.0662
22.6217 | -105.0546
22.6253 | -105.0445
22.629 | -105.0351
22.6327 | -105.0254
22.6363 | -105.0164
22.64 | -105.0064
22.6437 | -104.9957
22.6473 | -104.9864
22.651 | -104.9777
22.6546 | -104.9702
22.6583 | -104.9618
22.662 | -104.9527
22.6656 | -104.9436
22.6693 | -104.9377
22.673 | -104.9332
22.6766 | -104.9279
22.6803 | -104.9239
22.6839 | -104.9217
22.6876 | -104.9198
22.6913 | -104.9186
22.6949 | -104.9153
22.6986 | -104.9134
22.7023 | -104.9123
22.7059 | -104.9129
22.7096 | -104.9141
22.7132 | -104.9153
22.7169 | -104.9175
22.7206 | -104.9222
22.7242 | -104.9237
22.7279 | -104.9259
22.7316 | -104.9297
22.7352 | -104.9351
22.7389 | -104.9386
22.7425 | -104.9424
22.7462 | -104.9454
22.7499 | -104.9489
22.7535 | -104.9547
22.7572 | -104.9603
22.7609 | -104.9654
22.7645 | -104.9715
22.7682 | -104.9785
22.7719 | -104.9855
22.7755 | -104.9926
22.7792 | -104.9988
22.7828 | -105.0052
22.7865 | -105.0112
22.7902 | -105.0174
22.7938 | -105.0229
22.7975 | -105.0262
22.8012 | -105.0312
22.8048 | -105.0351
22.8085 | -105.0394
22.8121 | -105.044
22.8158 | -105.0475
22.8195 | -105.0501
22.8231 | -105.0525
22.8268 | -105.0551
22.8305 | -105.0569
22.8341 | -105.0585
22.8378 | -105.0609
22.8414 | -105.0633
22.8451 | -105.0649
22.8488 | -105.0678
22.8524 | -105.0715
22.8561 | -105.0752
22.8598 | -105.0768
22.8634 | -105.0739
22.8671 | -105.0711
22.8707 | -105.0642
22.8744 | -105.055
22.8781 | -105.0486
22.8817 | -105.0458
22.8854 | -105.0416
22.8891 | -105.037
22.8927 | -105.032
22.8964 | -105.0265
22.9001 | -105.0194
22.9037 | -105.012
22.9074 | -105.006
22.911 | -105.0011
22.9147 | -104.9949
22.9184 | -104.9869
22.922 | -104.9769
22.9257 | -104.9655
22.9294 | -104.953
22.933 | -104.9403
22.9367 | -104.9275
22.9403 | -104.9142
22.944 | -104.9
22.9477 | -104.8855
22.9513 | -104.8718
22.955 | -104.8588
22.9587 | -104.844
22.9623 | -104.8313
22.966 | -104.8182
22.9696 | -104.8052
22.9733 | -104.7908
22.977 | -104.7761
22.9806 | -104.7622
22.9843 | -104.7477
22.988 | -104.7353
22.9916 | -104.7231
22.9953 | -104.7094
22.9989 | -104.6976
23.0026 | -104.6872
23.0063 | -104.6771
23.0099 | -104.6674
23.0136 | -104.6579
23.0173 | -104.6499
23.0209 | -104.6429
23.0246 | -104.638
23.0282 | -104.634
23.0319 | -104.6295
23.0356 | -104.6276
23.0392 | -104.627
23.0429 | -104.6272
23.0466 | -104.6308
23.0502 | -104.6354
23.0539 | -104.6411
23.0576 | -104.6484
23.0612 | -104.6553
23.0649 | -104.6636
23.0685 | -104.6713
23.0722 | -104.6814
23.0759 | -104.6926
23.0795 | -104.7038
23.0832 | -104.7164
23.0869 | -104.7307
23.0905 | -104.7448
23.0942 | -104.7599
23.0978 | -104.7757
23.1015 | -104.7939
23.1052 | -104.8112
23.1088 | -104.8289
23.1125 | -104.8476
23.1162 | -104.8668
23.1198 | -104.8858
23.1235 | -104.9055
23.1271 | -104.9258
23.1308 | -104.9457
23.1345 | -104.9648
23.1381 | -104.9837
23.1418 | -105.002
23.1455 | -105.0209
23.1491 | -105.0405
23.1528 | -105.0595
23.1564 | -105.078
23.1601 | -105.0972
23.1638 | -105.1156
23.1674 | -105.1312
23.1711 | -105.1471
23.1748 | -105.164
23.1784 | -105.182
23.1821 | -105.2009
23.1858 | -105.2188
23.1894 | -105.2394
23.1931 | -105.2613
23.1967 | -105.2852
23.2004 | -105.3084
23.2041 | -105.329
23.2077 | -105.3455
23.2114 | -105.359
23.2151 | -105.3712
23.2187 | -105.3837
23.2224 | -105.3962
23.226 | -105.4088
23.2297 | -105.4223
23.2334 | -105.4359
23.237 | -105.4484
23.2407 | -105.4618
23.2444 | -105.4763
23.248 | -105.4864
23.2517 | -105.4999
23.2553 | -105.5122
23.259 | -105.5245
23.2627 | -105.5377
23.2663 | -105.5509
23.27 | -105.5598
23.2737 | -105.5677
23.2773 | -105.5771
23.281 | -105.5843
23.2846 | -105.5904
23.2883 | -105.5982
23.292 | -105.6032
23.2956 | -105.6076
23.2993 | -105.6113
23.303 | -105.6164
23.3066 | -105.621
23.3103 | -105.6232
23.314 | -105.6252
23.3176 | -105.6259
23.3213 | -105.6239
23.3249 | -105.6225
23.3286 | -105.6205
23.3323 | -105.618
23.3359 | -105.6154
23.3396 | -105.6132
23.3433 | -105.6107
23.3469 | -105.6081
23.3506 | -105.6043
23.3542 | -105.6
23.3579 | -105.5946
23.3616 | -105.5908
23.3652 | -105.586
23.3689 | -105.5811
23.3726 | -105.5768
23.3762 | -105.5712
23.3799 | -105.564
23.3835 | -105.5573
23.3872 | -105.5511
23.3909 | -105.547
23.3945 | -105.5436
23.3982 | -105.5406
23.4019 | -105.5376
23.4055 | -105.535
23.4092 | -105.5329
23.4128 | -105.5313
23.4165 | -105.5301
23.4202 | -105.5299
23.4238 | -105.5296
23.4275 | -105.5287
23.4312 | -105.528
23.4348 | -105.5298
23.4385 | -105.5321
23.4422 | -105.534
23.4458 | -105.5367
23.4495 | -105.5394
23.4531 | -105.5424
23.4568 | -105.5449
23.4605 | -105.5501
23.4641 | -105.5578
23.4678 | -105.5656
23.4715 | -105.5739
23.4751 | -105.5812
23.4788 | -105.5879
23.4824 | -105.5966
23.4861 | -105.6062
23.4898 | -105.6142
23.4934 | -105.6234
23.4971 | -105.6364
23.5008 | -105.6481
23.5044 | -105.6574
23.5081 | -105.6656
23.5117 | -105.6734
23.5154 | -105.6836
23.5191 | -105.6938
23.5227 | -105.7034
23.5264 | -105.7131
23.5301 | -105.7247
23.5337 | -105.7365
23.5374 | -105.7452
23.541 | -105.7564
23.5447 | -105.769
23.5484 | -105.781
23.552 | -105.7931
23.5557 | -105.8051
23.5594 | -105.8149
23.563 | -105.8232
23.5667 | -105.8326
23.5703 | -105.8425
23.574 | -105.8492
23.5777 | -105.8563
23.5813 | -105.864
23.585 | -105.871
23.5887 | -105.8787
23.5923 | -105.8851
23.596 | -105.8892
23.5997 | -105.8931
23.6033 | -105.899
23.607 | -105.9055
23.6106 | -105.9101
23.6143 | -105.9152
23.618 | -105.9202
23.6216 | -105.9235
23.6253 | -105.9265
23.629 | -105.93
23.6326 | -105.9347
23.6363 | -105.9385
23.6399 | -105.9413
23.6436 | -105.9452
23.6473 | -105.9498
23.6509 | -105.952
23.6546 | -105.9562
23.6583 | -105.9612
23.6619 | -105.9667
23.6656 | -105.9688
23.6692 | -105.9709
23.6729 | -105.9726
23.6766 | -105.9729
23.6802 | -105.9729
23.6839 | -105.9749
23.6876 | -105.9777
23.6912 | -105.98
23.6949 | -105.9814
23.6985 | -105.9829
23.7022 | -105.984
23.7059 | -105.9854
23.7095 | -105.9862
23.7132 | -105.9869
23.7169 | -105.9881
23.7205 | -105.9888
23.7242 | -105.9874
23.7279 | -105.9853
23.7315 | -105.9851
23.7352 | -105.983
23.7388 | -105.9812
23.7425 | -105.9779
23.7462 | -105.9752
23.7498 | -105.9722
23.7535 | -105.9694
23.7572 | -105.9672
23.7608 | -105.9664
23.7645 | -105.9651
23.7681 | -105.9639
23.7718 | -105.9625
23.7755 | -105.96
23.7791 | -105.9583
23.7828 | -105.9561
23.7865 | -105.955
23.7901 | -105.9561
23.7938 | -105.9576
23.7974 | -105.9563
23.8011 | -105.9528
23.8048 | -105.9505
23.8084 | -105.9502
23.8121 | -105.9505
23.8158 | -105.9524
23.8194 | -105.9517
23.8231 | -105.9504
23.8267 | -105.9511
23.8304 | -105.9538
23.8341 | -105.9543
23.8377 | -105.955
23.8414 | -105.9585
23.8451 | -105.9613
23.8487 | -105.9628
23.8524 | -105.9648
23.8561 | -105.9686
23.8597 | -105.9719
23.8634 | -105.974
23.867 | -105.9755
23.8707 | -105.977
23.8744 | -105.9825
23.878 | -105.9852
23.8817 | -105.9851
23.8854 | -105.9862
23.889 | -105.9854
23.8927 | -105.9844
23.8963 | -105.9869
23.9 | -105.989
23.9037 | -105.9907
23.9073 | -105.9917
23.911 | -105.9923
23.9147 | -105.9926
23.9183 | -105.9918
23.922 | -105.9913
23.9256 | -105.9913
23.9293 | -105.9928
23.933 | -105.9923
23.9366 | -105.9918
23.9403 | -105.9918
23.944 | -105.9904
23.9476 | -105.9895
23.9513 | -105.9879
23.9549 | -105.9863
23.9586 | -105.9856
23.9623 | -105.9862
23.9659 | -105.9849
23.9696 | -105.9807
23.9733 | -105.976
23.9769 | -105.9741
23.9806 | -105.9735
23.9842 | -105.9732
23.9879 | -105.975
23.9916 | -105.9746
23.9952 | -105.9748
23.9989 | -105.9752
24.0026 | -105.9775
24.0062 | -105.9806
24.0099 | -105.9844
24.0136 | -105.9891
24.0172 | -105.995
24.0209 | -105.999
24.0245 | -106.0023
24.0282 | -106.0062
24.0319 | -106.0113
24.0355 | -106.0163
24.0392 | -106.0226
24.0429 | -106.0308
24.0465 | -106.0383
24.0502 | -106.0467
24.0538 | -106.0552
24.0575 | -106.0628
24.0612 | -106.0716
24.0648 | -106.0837
24.0685 | -106.094
24.0722 | -106.1054
24.0758 | -106.1155
24.0795 | -106.1251
24.0831 | -106.1353
24.0868 | -106.1448
24.0905 | -106.1537
24.0941 | -106.1627
24.0978 | -106.1719
24.1015 | -106.1806
24.1051 | -106.1882
24.1088 | -106.1962
24.1124 | -106.2033
24.1161 | -106.2117
24.1198 | -106.2184
24.1234 | -106.2243
24.1271 | -106.2287
24.1308 | -106.2321
24.1344 | -106.2348
24.1381 | -106.2385
24.1418 | -106.243
24.1454 | -106.2461
24.1491 | -106.2463
24.1527 | -106.249
24.1564 | -106.25
24.1601 | -106.2505
24.1637 | -106.2483
24.1674 | -106.2466
24.1711 | -106.243
24.1747 | -106.2406
24.1784 | -106.2373
24.182 | -106.2295
24.1857 | -106.2243
24.1894 | -106.2217
24.193 | -106.2184
24.1967 | -106.2129
24.2004 | -106.2066
24.204 | -106.1995
24.2077 | -106.1932
24.2113 | -106.1865
24.215 | -106.1787
24.2187 | -106.1739
24.2223 | -106.1699
24.226 | -106.1644
24.2297 | -106.1594
24.2333 | -106.1557
24.237 | -106.149
24.2406 | -106.1433
24.2443 | -106.1383
24.248 | -106.1332
24.2516 | -106.1296
24.2553 | -106.1252
24.259 | -106.1202
24.2626 | -106.114
24.2663 | -106.1062
24.27 | -106.099
24.2736 | -106.093
24.2773 | -106.0864
24.2809 | -106.0779
24.2846 | -106.0701
24.2883 | -106.0615
24.2919 | -106.0533
24.2956 | -106.0454
24.2993 | -106.0379
24.3029 | -106.0281
24.3066 | -106.0178
24.3102 | -106.0095
24.3139 | -106.002
24.3176 | -105.9952
24.3212 | -105.9864
24.3249 | -105.9749
24.3286 | -105.9631
24.3322 | -105.951
24.3359 | -105.941
24.3395 | -105.9316
24.3432 | -105.9226
24.3469 | -105.9139
24.3505 | -105.9046
24.3542 | -105.8957
24.3579 | -105.8875
24.3615 | -105.878
24.3652 | -105.8701
24.3688 | -105.8631
24.3725 | -105.8557
24.3762 | -105.845
24.3798 | -105.8349
24.3835 | -105.8265
24.3872 | -105.8174
24.3908 | -105.8079
24.3945 | -105.7984
24.3982 | -105.7888
24.4018 | -105.7813
24.4055 | -105.7724
24.4091 | -105.7636
24.4128 | -105.7564
24.4165 | -105.749
24.4201 | -105.741
24.4238 | -105.7345
24.4275 | -105.7263
24.4311 | -105.7184
24.4348 | -105.712
24.4384 | -105.7044
24.4421 | -105.6943
24.4458 | -105.6864
24.4494 | -105.6795
24.4531 | -105.6725
24.4568 | -105.667
24.4604 | -105.6606
24.4641 | -105.6536
24.4677 | -105.6459
24.4714 | -105.6384
24.4751 | -105.6299
24.4787 | -105.6213
24.4824 | -105.6131
24.4861 | -105.6047
24.4897 | -105.5971
24.4934 | -105.5917
24.497 | -105.5844
24.5007 | -105.5773
24.5044 | -105.5705
24.508 | -105.5649
24.5117 | -105.5562
24.5154 | -105.5479
24.519 | -105.5408
24.5227 | -105.5352
24.5264 | -105.5303
24.53 | -105.5238
24.5337 | -105.5156
24.5373 | -105.5066
24.541 | -105.4981
24.5447 | -105.4911
24.5483 | -105.484
24.552 | -105.4755
24.5557 | -105.4668
24.5593 | -105.4584
24.563 | -105.4489
24.5666 | -105.4374
24.5703 | -105.4225
24.574 | -105.4061
24.5776 | -105.3896
24.5813 | -105.3726
24.585 | -105.3536
24.5886 | -105.3235
24.5923 | -105.2807
24.5959 | -105.2369
24.5996 | -105.1973
24.6033 | -105.1627
24.6069 | -105.1228
24.6106 | -105.0786
24.6143 | -105.0332
24.6179 | -104.9793
24.6216 | -104.9286
24.6252 | -104.8786
24.6289 | -104.8245
24.6326 | -104.7705
24.6362 | -104.7143
24.6399 | -104.6573
24.6436 | -104.5982
24.6472 | -104.5384
24.6509 | -104.4766
24.6545 | -104.4116
24.6582 | -104.3442
24.6619 | -104.2755
24.6655 | -104.2057
24.6692 | -104.1348
24.6729 | -104.0631
24.6765 | -103.9895
24.6802 | -103.9154
24.6839 | -103.8416
24.6875 | -103.767
24.6912 | -103.6913
24.6948 | -103.6153
24.6985 | -103.5412
24.7022 | -103.4682
24.7058 | -103.3964
24.7095 | -103.327
24.7132 | -103.2596
24.7168 | -103.1945
24.7205 | -103.1322
24.7241 | -103.073
24.7278 | -103.017
24.7315 | -102.9633
24.7351 | -102.9127
24.7388 | -102.8655
24.7425 | -102.8228
24.7461 | -102.7832
24.7498 | -102.7468
24.7534 | -102.7149
24.7571 | -102.6864
24.7608 | -102.6629
24.7644 | -102.6447
24.7681 | -102.6309
24.7718 | -102.6211
24.7754 | -102.6168
24.7791 | -102.6165
24.7827 | -102.6209
24.7864 | -102.6301
24.7901 | -102.6445
24.7937 | -102.6634
24.7974 | -102.6857
24.8011 | -102.7114
24.8047 | -102.7412
24.8084 | -102.7758
24.8121 | -102.8146
24.8157 | -102.8584
24.8194 | -102.9055
24.823 | -102.9559
24.8267 | -103.0102
24.8304 | -103.067
24.834 | -103.1415
24.8377 | -103.2011
24.8414 | -103.2626
24.845 | -103.326
24.8487 | -103.3912
24.8523 | -103.4582
24.856 | -103.5267
24.8597 | -103.5969
24.8633 | -103.6683
24.867 | -103.739
24.8707 | -103.8101
24.8743 | -103.8808
24.878 | -103.9524
24.8816 | -104.0239
24.8853 | -104.0954
24.889 | -104.1657
24.8926 | -104.2341
24.8963 | -104.3013
24.9 | -104.3681
24.9036 | -104.433
24.9073 | -104.4957
24.9109 | -104.5574
24.9146 | -104.6171
24.9183 | -104.6756
24.9219 | -104.734
24.9256 | -104.7909
24.9293 | -104.8422
24.9329 | -104.8911
24.9366 | -104.9377
24.9402 | -104.9903
24.9439 | -105.0482
24.9476 | -105.1005
24.9512 | -105.1433
24.9549 | -105.1762
24.9586 | -105.2092
24.9622 | -105.2405
24.9659 | -105.2675
24.9696 | -105.2981
24.9732 | -105.3195
24.9769 | -105.3359
24.9805 | -105.3511
24.9842 | -105.3621
24.9879 | -105.3708
24.9915 | -105.3763
24.9952 | -105.3821
24.9989 | -105.3855
25.0025 | -105.3873
25.0062 | -105.3902
25.0098 | -105.3936
25.0135 | -105.3962
25.0172 | -105.3958
25.0208 | -105.3955
25.0245 | -105.3972
25.0282 | -105.3996
25.0318 | -105.3998
25.0355 | -105.3993
25.0391 | -105.3983
25.0428 | -105.3982
25.0465 | -105.399
25.0501 | -105.3989
25.0538 | -105.3974
25.0575 | -105.3959
25.0611 | -105.3939
25.0648 | -105.3898
25.0684 | -105.3863
25.0721 | -105.3822
25.0758 | -105.3777
25.0794 | -105.3738
25.0831 | -105.3704
25.0868 | -105.3673
25.0904 | -105.3646
25.0941 | -105.3602
25.0978 | -105.3572
25.1014 | -105.3541
25.1051 | -105.3514
25.1087 | -105.3497
25.1124 | -105.3461
25.1161 | -105.3423
25.1197 | -105.3396
25.1234 | -105.3359
25.1271 | -105.3318
25.1307 | -105.3293
25.1344 | -105.3258
25.138 | -105.3208
25.1417 | -105.3175
25.1454 | -105.3133
25.149 | -105.3116
25.1527 | -105.3104
25.1564 | -105.3084
25.16 | -105.3052
25.1637 | -105.3011
25.1673 | -105.2961
25.171 | -105.2935
25.1747 | -105.2907
25.1783 | -105.287
25.182 | -105.2838
25.1857 | -105.2816
25.1893 | -105.2797
25.193 | -105.2775
25.1966 | -105.2755
25.2003 | -105.2741
25.204 | -105.2741
25.2076 | -105.2739
25.2113 | -105.2737
25.215 | -105.2747
25.2186 | -105.2758
25.2223 | -105.2765
25.226 | -105.2761
25.2296 | -105.2775
25.2333 | -105.2799
25.2369 | -105.2827
25.2406 | -105.2846
25.2443 | -105.2887
25.2479 | -105.2927
25.2516 | -105.2963
25.2553 | -105.2995
25.2589 | -105.3032
25.2626 | -105.3092
25.2662 | -105.3143
25.2699 | -105.3197
25.2736 | -105.3241
25.2772 | -105.3304
25.2809 | -105.3373
25.2846 | -105.3447
25.2882 | -105.353
25.2919 | -105.3611
25.2955 | -105.3676
25.2992 | -105.3745
25.3029 | -105.3806
25.3065 | -105.3849
25.3102 | -105.3909
25.3139 | -105.3956
25.3175 | -105.3996
25.3212 | -105.4045
25.3248 | -105.41
25.3285 | -105.4157
25.3322 | -105.4213
25.3358 | -105.4264
25.3395 | -105.4293
25.3432 | -105.432
25.3468 | -105.4333
25.3505 | -105.433
25.3542 | -105.4341
25.3578 | -105.4346
25.3615 | -105.4259
25.3651 | -105.4245
25.3688 | -105.4224
25.3725 | -105.4199
25.3761 | -105.4193
25.3798 | -105.4168
25.3835 | -105.4124
25.3871 | -105.4067
25.3908 | -105.4014
25.3944 | -105.3937
25.3981 | -105.3857
25.4018 | -105.3769
25.4054 | -105.3676
25.4091 | -105.36
25.4128 | -105.3528
25.4164 | -105.3444
25.4201 | -105.3348
25.4237 | -105.3269
25.4274 | -105.3178
25.4311 | -105.3075
25.4347 | -105.2985
25.4384 | -105.2864
25.4421 | -105.2731
25.4457 | -105.2617
25.4494 | -105.2507
25.453 | -105.2399
25.4567 | -105.2301
25.4604 | -105.22
25.464 | -105.2093
25.4677 | -105.1986
25.4714 | -105.1894
25.475 | -105.1803
25.4787 | -105.1713
25.4824 | -105.1639
25.486 | -105.1563
25.4897 | -105.1474
25.4933 | -105.1391
25.497 | -105.1332
25.5007 | -105.1263
25.5043 | -105.1204
25.508 | -105.1146
25.5117 | -105.1092
25.5153 | -105.1058
25.519 | -105.1027
25.5226 | -105.1013
25.5263 | -105.1022
25.53 | -105.1025
25.5336 | -105.1045
25.5373 | -105.1065
25.541 | -105.1095
25.5446 | -105.1133
25.5483 | -105.1185
25.5519 | -105.1239
25.5556 | -105.1296
25.5593 | -105.134
25.5629 | -105.1392
25.5666 | -105.1472
25.5703 | -105.1563
25.5739 | -105.1646
25.5776 | -105.1729
25.5812 | -105.1831
25.5849 | -105.1941
25.5886 | -105.2047
25.5922 | -105.2152
25.5959 | -105.225
25.5996 | -105.2373
25.6032 | -105.2485
25.6069 | -105.2609
25.6105 | -105.2738
25.6142 | -105.287
25.6179 | -105.2986
25.6215 | -105.3078
25.6252 | -105.3183
25.6289 | -105.3296
25.6325 | -105.3416
25.6362 | -105.3519
25.6399 | -105.3627
25.6435 | -105.3739
25.6472 | -105.3833
25.6508 | -105.3928
25.6545 | -105.4024
25.6582 | -105.411
25.6618 | -105.4202
25.6655 | -105.4302
25.6692 | -105.4378
25.6728 | -105.4466
25.6765 | -105.4568
25.6801 | -105.4637
25.6838 | -105.4691
25.6875 | -105.4742
25.6911 | -105.478
25.6948 | -105.4832
25.6985 | -105.4906
25.7021 | -105.4971
25.7058 | -105.5034
25.7094 | -105.5095
25.7131 | -105.5142
25.7168 | -105.5206
25.7204 | -105.5263
25.7241 | -105.5306
25.7278 | -105.5349
25.7314 | -105.5387
25.7351 | -105.5423
25.7387 | -105.5445
25.7424 | -105.5471
25.7461 | -105.5491
25.7497 | -105.5524
25.7534 | -105.5572
25.7571 | -105.5621
25.7607 | -105.5665
25.7644 | -105.5708
25.7681 | -105.5753
25.7717 | -105.5769
25.7754 | -105.579
25.779 | -105.5819
25.7827 | -105.5853
25.7864 | -105.5885
25.79 | -105.5926
25.7937 | -105.597
25.7974 | -105.6011
25.801 | -105.6072
25.8047 | -105.613
25.8083 | -105.6172
25.812 | -105.6225
25.8157 | -105.6268
25.8193 | -105.6309
25.823 | -105.6347
25.8267 | -105.6377
25.8303 | -105.6415
25.834 | -105.6457
25.8376 | -105.6495
25.8413 | -105.6538
25.845 | -105.6582
25.8486 | -105.6619
25.8523 | -105.6685
25.856 | -105.6749
25.8596 | -105.6797
25.8633 | -105.6832
25.8669 | -105.6861
25.8706 | -105.6892
25.8743 | -105.6935
25.8779 | -105.6969
25.8816 | -105.7006
25.8853 | -105.7041
25.8889 | -105.7051
25.8926 | -105.7067
25.8962 | -105.7086
25.8999 | -105.7105
25.9036 | -105.714
25.9072 | -105.7186
25.9109 | -105.7228
25.9146 | -105.7258
25.9182 | -105.7293
25.9219 | -105.7347
25.9256 | -105.7423
25.9292 | -105.7482
25.9329 | -105.7519
25.9365 | -105.7575
25.9402 | -105.7632
25.9439 | -105.77
25.9475 | -105.7769
25.9512 | -105.7828
25.9549 | -105.7885
25.9585 | -105.796
25.9622 | -105.8039
25.9658 | -105.8109
25.9695 | -105.8185
25.9732 | -105.8266
25.9768 | -105.8349
25.9805 | -105.8427
25.9842 | -105.8512
25.9878 | -105.861
25.9915 | -105.8685
25.9951 | -105.8763
25.9988 | -105.8854
26.0025 | -105.8941
26.0061 | -105.9028
26.0098 | -105.9099
26.0135 | -105.9169
26.0171 | -105.9272
26.0208 | -105.9372
26.0244 | -105.9447
26.0281 | -105.9525
26.0318 | -105.96
26.0354 | -105.9672
26.0391 | -105.9742
26.0428 | -105.9811
26.0464 | -105.9888
26.0501 | -105.9953
26.0538 | -106.002
26.0574 | -106.0099
26.0611 | -106.0167
26.0647 | -106.0203
26.0684 | -106.0241
26.0721 | -106.0291
26.0757 | -106.0317
26.0794 | -106.035
26.0831 | -106.0369
26.0867 | -106.0396
26.0904 | -106.0419
26.094 | -106.0435
26.0977 | -106.0449
26.1014 | -106.0467
26.105 | -106.0477
26.1087 | -106.0496
26.1124 | -106.0518
26.116 | -106.0521
26.1197 | -106.0508
26.1233 | -106.0496
26.127 | -106.0475
26.1307 | -106.0451
26.1343 | -106.0429
26.138 | -106.0395
26.1417 | -106.0341
26.1453 | -106.03
26.149 | -106.0266
26.1526 | -106.0217
26.1563 | -106.0163
26.16 | -106.0127
26.1636 | -106.0083
26.1673 | -106.0037
26.171 | -105.9981
26.1746 | -105.9939
26.1783 | -105.9895
26.182 | -105.9848
26.1856 | -105.9784
26.1893 | -105.9723
26.1929 | -105.9667
26.1966 | -105.962
26.2003 | -105.9567
26.2039 | -105.9527
26.2076 | -105.9478
26.2113 | -105.9431
26.2149 | -105.9389
26.2186 | -105.9337
26.2222 | -105.929
26.2259 | -105.9258
26.2296 | -105.923
26.2332 | -105.919
26.2369 | -105.9149
26.2406 | -105.9123
26.2442 | -105.9108
26.2479 | -105.9094
26.2515 | -105.9077
26.2552 | -105.9065
26.2589 | -105.9056
26.2625 | -105.9047
26.2662 | -105.9037
26.2699 | -105.9028
26.2735 | -105.9014
26.2772 | -105.8995
26.2808 | -105.8972
26.2845 | -105.8958
26.2882 | -105.8946
26.2918 | -105.8934
26.2955 | -105.8916
26.2992 | -105.8924
26.3028 | -105.8947
26.3065 | -105.8942
26.3102 | -105.8948
26.3138 | -105.8957
26.3175 | -105.8961
26.3211 | -105.8976
26.3248 | -105.8977
26.3285 | -105.8965
26.3321 | -105.896
26.3358 | -105.8964
26.3395 | -105.8976
26.3431 | -105.8997
26.3468 | -105.9011
26.3504 | -105.9028
26.3541 | -105.9038
26.3578 | -105.9059
26.3614 | -105.9092
26.3651 | -105.9084
26.3688 | -105.9082
26.3724 | -105.9102
26.3761 | -105.9106
26.3797 | -105.9123
26.3834 | -105.9141
26.3871 | -105.9138
26.3907 | -105.9144
26.3944 | -105.9175
26.3981 | -105.92
26.4017 | -105.9219
26.4054 | -105.9249
26.409 | -105.9276
26.4127 | -105.9288
26.4164 | -105.9293
26.42 | -105.9314
26.4237 | -105.9336
26.4274 | -105.9344
26.431 | -105.9351
26.4347 | -105.9363
26.4383 | -105.9396
26.442 | -105.9437
26.4457 | -105.9459
26.4493 | -105.9495
26.453 | -105.9521
26.4567 | -105.9535
26.4603 | -105.9546
26.464 | -105.9563
26.4677 | -105.9573
26.4713 | -105.9586
26.475 | -105.9604
26.4786 | -105.9631
26.4823 | -105.964
26.486 | -105.9649
26.4896 | -105.9663
26.4933 | -105.9679
26.497 | -105.9702
26.5006 | -105.9736
26.5043 | -105.976
26.5079 | -105.9767
26.5116 | -105.9788
26.5153 | -105.982
26.5189 | -105.9835
26.5226 | -105.9839
26.5263 | -105.9853
26.5299 | -105.9867
26.5336 | -105.9872
26.5372 | -105.9865
26.5409 | -105.9875
26.5446 | -105.9887
26.5482 | -105.9866
26.5519 | -105.9854
26.5556 | -105.9871
26.5592 | -105.9884
26.5629 | -105.9899
26.5665 | -105.9939
26.5702 | -105.9951
26.5739 | -105.9957
26.5775 | -105.9949
26.5812 | -105.9956
26.5849 | -105.9958
26.5885 | -105.9959
26.5922 | -105.9968
26.5959 | -105.997
26.5995 | -105.997
26.6032 | -105.9975
26.6068 | -105.9969
26.6105 | -105.9971
26.6142 | -105.995
26.6178 | -105.9928
26.6215 | -105.9912
26.6252 | -105.9886
26.6288 | -105.9874
26.6325 | -105.9858
26.6361 | -105.9838
26.6398 | -105.982
26.6435 | -105.9795
26.6471 | -105.9791
26.6508 | -105.979
26.6545 | -105.9777
26.6581 | -105.977
26.6618 | -105.9758
26.6654 | -105.9729
26.6691 | -105.9705
26.6728 | -105.966
26.6764 | -105.963
26.6801 | -105.9615
26.6838 | -105.9585
26.6874 | -105.9567
26.6911 | -105.9553
26.6947 | -105.9531
26.6984 | -105.9519
26.7021 | -105.9503
26.7057 | -105.9503
26.7094 | -105.9481
26.7131 | -105.9444
26.7167 | -105.9409
26.7204 | -105.9397
26.7241 | -105.9389
26.7277 | -105.939
26.7314 | -105.9365
26.735 | -105.9345
26.7387 | -105.9347
26.7424 | -105.9325
26.746 | -105.9289
26.7497 | -105.9268
26.7534 | -105.927
26.757 | -105.927
26.7607 | -105.9251
26.7643 | -105.923
26.768 | -105.921
26.7717 | -105.9189
26.7753 | -105.9182
26.779 | -105.9173
26.7827 | -105.9168
26.7863 | -105.9171
26.79 | -105.9179
26.7936 | -105.9181
26.7973 | -105.9181
26.801 | -105.9178
26.8046 | -105.9192
26.8083 | -105.9203
26.812 | -105.9204
26.8156 | -105.9205
26.8193 | -105.9202
26.8229 | -105.9202
26.8266 | -105.9199
26.8303 | -105.9212
26.8339 | -105.9239
26.8376 | -105.9257
26.8413 | -105.9278
26.8449 | -105.9314
26.8486 | -105.9332
26.8523 | -105.9351
26.8559 | -105.9368
26.8596 | -105.9396
26.8632 | -105.9422
26.8669 | -105.9449
26.8706 | -105.9473
26.8742 | -105.9512
26.8779 | -105.9556
26.8816 | -105.9604
26.8852 | -105.9639
26.8889 | -105.9679
26.8925 | -105.9736
26.8962 | -105.979
26.8999 | -105.984
26.9035 | -105.9905
26.9072 | -105.9971
26.9109 | -106.0033
26.9145 | -106.0106
26.9182 | -106.0165
26.9218 | -106.0229
26.9255 | -106.031
26.9292 | -106.0377
26.9328 | -106.0444
26.9365 | -106.0525
26.9402 | -106.0595
26.9438 | -106.0637
26.9475 | -106.0681
26.9511 | -106.0733
26.9548 | -106.0798
26.9585 | -106.0864
26.9621 | -106.0918
26.9658 | -106.0982
26.9695 | -106.1051
26.9731 | -106.1102
26.9768 | -106.1154
26.9805 | -106.1193
26.9841 | -106.125
26.9878 | -106.1316
26.9914 | -106.1356
26.9951 | -106.1393
26.9988 | -106.1432
27.0024 | -106.1455
27.0061 | -106.1476
27.0098 | -106.1502
27.0134 | -106.152
27.0171 | -106.1528
27.0207 | -106.154
27.0244 | -106.1552
27.0281 | -106.1547
27.0317 | -106.1536
27.0354 | -106.1528
27.0391 | -106.1511
27.0427 | -106.1496
27.0464 | -106.1477
27.05 | -106.1448
27.0537 | -106.1417
27.0574 | -106.1361
27.061 | -106.1298
27.0647 | -106.125
27.0684 | -106.1188
27.072 | -106.112
27.0757 | -106.1063
27.0793 | -106.1007
27.083 | -106.0956
27.0867 | -106.0887
27.0903 | -106.0821
27.094 | -106.0761
27.0977 | -106.068
27.1013 | -106.06
27.105 | -106.0532
27.1086 | -106.0462
27.1123 | -106.0378
27.116 | -106.0307
27.1196 | -106.0248
27.1233 | -106.0172
27.127 | -106.0064
27.1306 | -105.9966
27.1343 | -105.9847
27.138 | -105.975
27.1416 | -105.9659
27.1453 | -105.9577
27.1489 | -105.9488
27.1526 | -105.9406
27.1563 | -105.9302
27.1599 | -105.9205
27.1636 | -105.9112
27.1673 | -105.9034
27.1709 | -105.8954
27.1746 | -105.8876
27.1782 | -105.878
27.1819 | -105.8683
27.1856 | -105.86
27.1892 | -105.8522
27.1929 | -105.8437
27.1966 | -105.8344
27.2002 | -105.8261
27.2039 | -105.8184
27.2075 | -105.8097
27.2112 | -105.8004
27.2149 | -105.7902
27.2185 | -105.7805
27.2222 | -105.7718
27.2259 | -105.7634
27.2295 | -105.756
27.2332 | -105.7484
27.2368 | -105.7411
27.2405 | -105.733
27.2442 | -105.7259
27.2478 | -105.7197
27.2515 | -105.7123
27.2552 | -105.7035
27.2588 | -105.6943
27.2625 | -105.6868
27.2662 | -105.6816
27.2698 | -105.6763
27.2735 | -105.6663
27.2771 | -105.6594
27.2808 | -105.6543
27.2845 | -105.6482
27.2881 | -105.6423
27.2918 | -105.6366
27.2955 | -105.6311
27.2991 | -105.6265
27.3028 | -105.621
27.3064 | -105.6155
27.3101 | -105.6115
27.3138 | -105.6066
27.3174 | -105.6007
27.3211 | -105.5959
27.3248 | -105.5928
27.3284 | -105.5911
27.3321 | -105.5905
27.3357 | -105.5883
27.3394 | -105.5849
27.3431 | -105.5809
27.3467 | -105.5778
27.3504 | -105.5737
27.3541 | -105.57
27.3577 | -105.5664
27.3614 | -105.5642
27.365 | -105.5636
27.3687 | -105.5607
27.3724 | -105.5586
27.376 | -105.5584
27.3797 | -105.5592
27.3834 | -105.5594
27.387 | -105.5594
27.3907 | -105.5597
27.3943 | -105.5577
27.398 | -105.5578
27.4017 | -105.5588
27.4053 | -105.5585
27.409 | -105.5588
27.4127 | -105.5586
27.4163 | -105.5596
27.42 | -105.5622
27.4237 | -105.5628
27.4273 | -105.5628
27.431 | -105.5648
27.4346 | -105.567
27.4383 | -105.5701
27.442 | -105.5727
27.4456 | -105.5752
27.4493 | -105.5772
27.453 | -105.5789
27.4566 | -105.5814
27.4603 | -105.5835
27.4639 | -105.5846
27.4676 | -105.5876
27.4713 | -105.5914
27.4749 | -105.5982
27.4786 | -105.6026
27.4823 | -105.6071
27.4859 | -105.6131
27.4896 | -105.6179
27.4932 | -105.624
27.4969 | -105.6306
27.5006 | -105.6381
27.5042 | -105.6442
27.5079 | -105.6502
27.5116 | -105.6592
27.5152 | -105.6695
27.5189 | -105.6776
27.5225 | -105.6857
27.5262 | -105.6912
27.5299 | -105.6993
27.5335 | -105.7071
27.5372 | -105.7148
27.5409 | -105.722
27.5445 | -105.7304
27.5482 | -105.7378
27.5519 | -105.7451
27.5555 | -105.7523
27.5592 | -105.7606
27.5628 | -105.7689
27.5665 | -105.7777
27.5702 | -105.7872
27.5738 | -105.7967
27.5775 | -105.8053
27.5812 | -105.8142
27.5848 | -105.8229
27.5885 | -105.8313
27.5921 | -105.8401
27.5958 | -105.8498
27.5995 | -105.8586
27.6031 | -105.868
27.6068 | -105.8775
27.6105 | -105.8858
27.6141 | -105.8931
27.6178 | -105.9004
27.6214 | -105.9071
27.6251 | -105.9144
27.6288 | -105.9228
27.6324 | -105.9307
27.6361 | -105.9365
27.6398 | -105.9443
27.6434 | -105.9514
27.6471 | -105.9594
27.6507 | -105.9672
27.6544 | -105.9741
27.6581 | -105.9797
27.6617 | -105.9874
27.6654 | -105.9916
27.6691 | -105.9942
27.6727 | -105.9976
27.6764 | -106.0021
27.6801 | -106.0056
27.6837 | -106.009
27.6874 | -106.0117
27.691 | -106.0157
27.6947 | -106.0199
27.6984 | -106.0231
27.702 | -106.0259
27.7057 | -106.029
27.7094 | -106.0327
27.713 | -106.0342
27.7167 | -106.034
27.7203 | -106.0311
27.724 | -106.0301
27.7277 | -106.0311
27.7313 | -106.0322
27.735 | -106.0325
27.7387 | -106.0322
27.7423 | -106.0318
27.746 | -106.0319
27.7496 | -106.03
27.7533 | -106.0262
27.757 | -106.0237
27.7606 | -106.0196
27.7643 | -106.0142
27.768 | -106.0086
27.7716 | -106.0035
27.7753 | -105.9984
27.7789 | -105.994
27.7826 | -105.9904
27.7863 | -105.9857
27.7899 | -105.9797
27.7936 | -105.9732
27.7973 | -105.9671
27.8009 | -105.9611
27.8046 | -105.954
27.8083 | -105.947
27.8119 | -105.9427
27.8156 | -105.9381
27.8192 | -105.9309
27.8229 | -105.9242
27.8266 | -105.9164
27.8302 | -105.9093
27.8339 | -105.9029
27.8376 | -105.8949
27.8412 | -105.8873
27.8449 | -105.8795
27.8485 | -105.8719
27.8522 | -105.8646
27.8559 | -105.8548
27.8595 | -105.8465
27.8632 | -105.8394
27.8669 | -105.8294
27.8705 | -105.8215
27.8742 | -105.813
27.8778 | -105.8035
27.8815 | -105.7947
27.8852 | -105.7858
27.8888 | -105.7778
27.8925 | -105.7683
27.8962 | -105.7585
27.8998 | -105.7497
27.9035 | -105.7403
27.9071 | -105.7318
27.9108 | -105.7234
27.9145 | -105.7147
27.9181 | -105.7061
27.9218 | -105.6959
27.9255 | -105.6871
27.9291 | -105.6787
27.9328 | -105.6702
27.9365 | -105.6621
27.9401 | -105.654
27.9438 | -105.6451
27.9474 | -105.6365
27.9511 | -105.6282
27.9548 | -105.62
27.9584 | -105.6123
27.9621 | -105.6045
27.9658 | -105.5962
27.9694 | -105.5888
27.9731 | -105.5797
27.9767 | -105.5706
27.9804 | -105.5625
27.9841 | -105.5554
27.9877 | -105.5486
27.9914 | -105.5403
27.9951 | -105.5316
27.9987 | -105.5248
28.0024 | -105.5159