## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 211 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 209
Hood | 02-92B/P | 149
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 994.1 | hPa |  | PRC PRO 01
Humidity | 43.7 | %/Hr
Internal temperature of calibration bench | 22.78 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 999 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.5 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.9 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3138 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 11/04/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 16/04/2018
Engineer | Anthony HUE | Approval |  | 23/04/2018
CEO | Robert GLEMOT | Approval |  | 23/04/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5728 | 57.5066
0.012209 | -36.1502 | 38.2683
0.018313 | -35.0876 | 28.157
0.024417 | -34.6303 | 21.9318
0.030521 | -34.5194 | 18.1804
0.036626 | -34.3767 | 15.129
0.04273 | -34.2813 | 12.7588
0.048834 | -34.2325 | 11.2443
0.054938 | -34.1646 | 10.1205
0.061043 | -34.1026 | 9.0384
0.067147 | -34.094 | 8.2094
0.073251 | -34.0965 | 7.4494
0.079355 | -34.1162 | 6.7926
0.08546 | -34.1309 | 6.2042
0.091564 | -34.126 | 5.7681
0.097668 | -34.1069 | 5.5242
0.10377 | -34.0839 | 5.2791
0.10988 | -34.0664 | 5.0011
0.11598 | -34.0577 | 4.7035
0.12209 | -34.0603 | 4.4141
0.12819 | -34.0709 | 4.1261
0.13429 | -34.0818 | 3.8352
0.1404 | -34.0843 | 3.5586
0.1465 | -34.0856 | 3.3027
0.15261 | -34.0822 | 3.0859
0.15871 | -34.0752 | 2.9271
0.16482 | -34.0669 | 2.8268
0.17092 | -34.0647 | 2.7892
0.17702 | -34.0626 | 2.7185
0.18313 | -34.0618 | 2.6419
0.18923 | -34.0615 | 2.5341
0.19534 | -34.0605 | 2.4085
0.20144 | -34.0594 | 2.2675
0.20754 | -34.0565 | 2.1476
0.21365 | -34.0555 | 2.0337
0.21975 | -34.0549 | 1.9363
0.22586 | -34.0558 | 1.849
0.23196 | -34.0581 | 1.7767
0.23807 | -34.0619 | 1.714
0.24417 | -34.0658 | 1.6558
0.25027 | -34.0696 | 1.5937
0.25638 | -34.0718 | 1.5318
0.26248 | -34.0736 | 1.4649
0.26859 | -34.0744 | 1.399
0.27469 | -34.0747 | 1.3268
0.2808 | -34.0757 | 1.2516
0.2869 | -34.0764 | 1.1796
0.293 | -34.0773 | 1.1117
0.29911 | -34.0781 | 1.0512
0.30521 | -34.0791 | 0.9921
0.31132 | -34.0798 | 0.95455
0.31742 | -34.081 | 0.91228
0.32353 | -34.0814 | 0.87835
0.32963 | -34.0823 | 0.85049
0.33573 | -34.0834 | 0.83155
0.34184 | -34.0848 | 0.81148
0.34794 | -34.087 | 0.78859
0.35405 | -34.0892 | 0.76551
0.36015 | -34.0918 | 0.74098
0.36626 | -34.0938 | 0.7141
0.37236 | -34.0959 | 0.67772
0.37846 | -34.0981 | 0.63353
0.38457 | -34.0989 | 0.57584
0.39067 | -34.0999 | 0.52893
0.39678 | -34.1005 | 0.47652
0.40288 | -34.1009 | 0.42198
0.40899 | -34.1007 | 0.36953
0.41509 | -34.1 | 0.31798
0.42119 | -34.0999 | 0.274
0.4273 | -34.0997 | 0.22707
0.4334 | -34.0996 | 0.18021
0.43951 | -34.0998 | 0.13618
0.44561 | -34.1006 | 0.095894
0.45172 | -34.1021 | 0.063739
0.45782 | -34.1029 | 0.026187
0.46392 | -34.1041 | -0.010317
0.47003 | -34.1052 | -0.04294
0.47613 | -34.106 | -0.077318
0.48224 | -34.1065 | -0.11233
0.48834 | -34.1056 | -0.14706
0.49445 | -34.1054 | -0.18238
0.50055 | -34.1048 | -0.21448
0.50665 | -34.1041 | -0.24868
0.51276 | -34.103 | -0.2827
0.51886 | -34.102 | -0.31661
0.52497 | -34.1007 | -0.34065
0.53107 | -34.0999 | -0.37012
0.53717 | -34.0991 | -0.39832
0.54328 | -34.0982 | -0.42103
0.54938 | -34.098 | -0.4435
0.55549 | -34.0979 | -0.46433
0.56159 | -34.099 | -0.47583
0.5677 | -34.0992 | -0.49034
0.5738 | -34.1002 | -0.50663
0.5799 | -34.101 | -0.52146
0.58601 | -34.102 | -0.53155
0.59211 | -34.1037 | -0.53777
0.59822 | -34.105 | -0.54726
0.60432 | -34.1065 | -0.55935
0.61043 | -34.1077 | -0.5673
0.61653 | -34.109 | -0.57765
0.62263 | -34.1105 | -0.58701
0.62874 | -34.1112 | -0.59786
0.63484 | -34.112 | -0.60729
0.64095 | -34.1125 | -0.61767
0.64705 | -34.1129 | -0.6276
0.65316 | -34.1131 | -0.64028
0.65926 | -34.1132 | -0.65632
0.66536 | -34.112 | -0.67409
0.67147 | -34.1116 | -0.68759
0.67757 | -34.1113 | -0.70217
0.68368 | -34.1105 | -0.72138
0.68978 | -34.1097 | -0.73891
0.69589 | -34.1087 | -0.75821
0.70199 | -34.1077 | -0.78124
0.70809 | -34.1068 | -0.80069
0.7142 | -34.106 | -0.82141
0.7203 | -34.1049 | -0.84088
0.72641 | -34.1042 | -0.86362
0.73251 | -34.1035 | -0.88386
0.73862 | -34.1037 | -0.90584
0.74472 | -34.1033 | -0.92607
0.75082 | -34.103 | -0.94843
0.75693 | -34.1033 | -0.97048
0.76303 | -34.1038 | -0.99271
0.76914 | -34.1056 | -1.0134
0.77524 | -34.1066 | -1.0332
0.78135 | -34.1078 | -1.0535
0.78745 | -34.1092 | -1.0737
0.79355 | -34.1104 | -1.0936
0.79966 | -34.1121 | -1.1133
0.80576 | -34.1144 | -1.1333
0.81187 | -34.116 | -1.1551
0.81797 | -34.1179 | -1.1766
0.82408 | -34.1196 | -1.1974
0.83018 | -34.1213 | -1.2182
0.83628 | -34.1232 | -1.2369
0.84239 | -34.1244 | -1.2577
0.84849 | -34.1258 | -1.2796
0.8546 | -34.1273 | -1.2997
0.8607 | -34.1285 | -1.3207
0.86681 | -34.1296 | -1.3398
0.87291 | -34.1306 | -1.3589
0.87901 | -34.1311 | -1.3771
0.88512 | -34.1318 | -1.3954
0.89122 | -34.1325 | -1.4139
0.89733 | -34.1329 | -1.4298
0.90343 | -34.1333 | -1.4476
0.90953 | -34.1335 | -1.4622
0.91564 | -34.1335 | -1.4782
0.92174 | -34.1337 | -1.4934
0.92785 | -34.1338 | -1.5082
0.93395 | -34.1338 | -1.5223
0.94006 | -34.134 | -1.5353
0.94616 | -34.1337 | -1.5473
0.95226 | -34.1338 | -1.5601
0.95837 | -34.1338 | -1.5741
0.96447 | -34.134 | -1.5872
0.97058 | -34.1338 | -1.6007
0.97668 | -34.1336 | -1.6142
0.98279 | -34.1333 | -1.6271
0.98889 | -34.1331 | -1.6379
0.99499 | -34.133 | -1.6513
1.0011 | -34.1328 | -1.6621
1.0072 | -34.1325 | -1.6738
1.0133 | -34.1321 | -1.6846
1.0194 | -34.1316 | -1.6976
1.0255 | -34.1313 | -1.7114
1.0316 | -34.1307 | -1.7234
1.0377 | -34.1303 | -1.7371
1.0438 | -34.1297 | -1.7499
1.0499 | -34.1289 | -1.7652
1.056 | -34.1282 | -1.7796
1.0621 | -34.1275 | -1.7939
1.0682 | -34.127 | -1.8064
1.0743 | -34.1264 | -1.8218
1.0805 | -34.1259 | -1.8363
1.0866 | -34.1249 | -1.8514
1.0927 | -34.124 | -1.8659
1.0988 | -34.1234 | -1.8808
1.1049 | -34.1226 | -1.895
1.111 | -34.1218 | -1.9104
1.1171 | -34.1211 | -1.925
1.1232 | -34.1206 | -1.9401
1.1293 | -34.1202 | -1.9542
1.1354 | -34.1199 | -1.9691
1.1415 | -34.1195 | -1.9847
1.1476 | -34.1188 | -2.0002
1.1537 | -34.1183 | -2.0149
1.1598 | -34.1184 | -2.0296
1.1659 | -34.1182 | -2.0429
1.172 | -34.1179 | -2.0568
1.1781 | -34.1177 | -2.0708
1.1842 | -34.1175 | -2.0849
1.1903 | -34.1179 | -2.0959
1.1964 | -34.1178 | -2.1093
1.2025 | -34.1179 | -2.1232
1.2086 | -34.1182 | -2.1356
1.2147 | -34.1182 | -2.1481
1.2209 | -34.1186 | -2.1593
1.227 | -34.1193 | -2.17
1.2331 | -34.1197 | -2.182
1.2392 | -34.1201 | -2.1939
1.2453 | -34.1205 | -2.2049
1.2514 | -34.1209 | -2.2148
1.2575 | -34.1215 | -2.2259
1.2636 | -34.1224 | -2.2349
1.2697 | -34.123 | -2.2442
1.2758 | -34.1236 | -2.2554
1.2819 | -34.1244 | -2.2654
1.288 | -34.125 | -2.2763
1.2941 | -34.1259 | -2.2848
1.3002 | -34.1266 | -2.2929
1.3063 | -34.1271 | -2.3021
1.3124 | -34.1278 | -2.3112
1.3185 | -34.1285 | -2.3204
1.3246 | -34.1292 | -2.3297
1.3307 | -34.1301 | -2.3384
1.3368 | -34.1309 | -2.3472
1.3429 | -34.1315 | -2.3563
1.349 | -34.132 | -2.3658
1.3551 | -34.1327 | -2.3754
1.3613 | -34.1333 | -2.3838
1.3674 | -34.1339 | -2.394
1.3735 | -34.1343 | -2.4039
1.3796 | -34.1348 | -2.4131
1.3857 | -34.1351 | -2.4226
1.3918 | -34.1356 | -2.4329
1.3979 | -34.1362 | -2.4422
1.404 | -34.1367 | -2.4537
1.4101 | -34.1371 | -2.4642
1.4162 | -34.1375 | -2.475
1.4223 | -34.138 | -2.4865
1.4284 | -34.1384 | -2.4967
1.4345 | -34.1387 | -2.5084
1.4406 | -34.139 | -2.522
1.4467 | -34.1394 | -2.5338
1.4528 | -34.1398 | -2.5453
1.4589 | -34.1403 | -2.5573
1.465 | -34.1406 | -2.5682
1.4711 | -34.141 | -2.5797
1.4772 | -34.1415 | -2.5932
1.4833 | -34.142 | -2.6063
1.4894 | -34.1425 | -2.6187
1.4955 | -34.143 | -2.6307
1.5016 | -34.1436 | -2.6434
1.5078 | -34.144 | -2.6574
1.5139 | -34.1444 | -2.6706
1.52 | -34.1448 | -2.6833
1.5261 | -34.1451 | -2.696
1.5322 | -34.1454 | -2.7102
1.5383 | -34.1458 | -2.7229
1.5444 | -34.1463 | -2.7356
1.5505 | -34.1468 | -2.7491
1.5566 | -34.1473 | -2.7623
1.5627 | -34.1477 | -2.7762
1.5688 | -34.1484 | -2.7899
1.5749 | -34.1487 | -2.8026
1.581 | -34.149 | -2.816
1.5871 | -34.1494 | -2.8299
1.5932 | -34.1497 | -2.8428
1.5993 | -34.1501 | -2.8572
1.6054 | -34.1505 | -2.8705
1.6115 | -34.1509 | -2.8847
1.6176 | -34.1512 | -2.8974
1.6237 | -34.1514 | -2.9111
1.6298 | -34.1517 | -2.9237
1.6359 | -34.152 | -2.9366
1.642 | -34.1524 | -2.9486
1.6482 | -34.1526 | -2.9616
1.6543 | -34.153 | -2.9754
1.6604 | -34.1533 | -2.9875
1.6665 | -34.1534 | -3
1.6726 | -34.1536 | -3.0126
1.6787 | -34.1539 | -3.0252
1.6848 | -34.1538 | -3.0379
1.6909 | -34.1541 | -3.0506
1.697 | -34.1543 | -3.0629
1.7031 | -34.1543 | -3.0756
1.7092 | -34.1543 | -3.0878
1.7153 | -34.1544 | -3.0998
1.7214 | -34.1542 | -3.1121
1.7275 | -34.1541 | -3.1248
1.7336 | -34.154 | -3.1378
1.7397 | -34.1542 | -3.1503
1.7458 | -34.1542 | -3.1625
1.7519 | -34.1542 | -3.1748
1.758 | -34.1542 | -3.1868
1.7641 | -34.1543 | -3.1985
1.7702 | -34.1543 | -3.2115
1.7763 | -34.1542 | -3.2231
1.7824 | -34.1542 | -3.2356
1.7885 | -34.1541 | -3.2482
1.7947 | -34.1542 | -3.2602
1.8008 | -34.1542 | -3.2716
1.8069 | -34.1542 | -3.2832
1.813 | -34.1543 | -3.2955
1.8191 | -34.1542 | -3.3083
1.8252 | -34.1543 | -3.3212
1.8313 | -34.1544 | -3.3336
1.8374 | -34.1545 | -3.3455
1.8435 | -34.1546 | -3.3576
1.8496 | -34.1546 | -3.3708
1.8557 | -34.1548 | -3.3836
1.8618 | -34.155 | -3.396
1.8679 | -34.1551 | -3.4082
1.874 | -34.1552 | -3.4203
1.8801 | -34.1553 | -3.4329
1.8862 | -34.1556 | -3.4449
1.8923 | -34.1557 | -3.4572
1.8984 | -34.1558 | -3.4695
1.9045 | -34.1558 | -3.4812
1.9106 | -34.1559 | -3.4936
1.9167 | -34.1559 | -3.5059
1.9228 | -34.156 | -3.5179
1.9289 | -34.1561 | -3.53
1.9351 | -34.1563 | -3.541
1.9412 | -34.1565 | -3.5523
1.9473 | -34.1568 | -3.5646
1.9534 | -34.1569 | -3.5761
1.9595 | -34.1571 | -3.5881
1.9656 | -34.1574 | -3.6006
1.9717 | -34.1576 | -3.6125
1.9778 | -34.1579 | -3.624
1.9839 | -34.1581 | -3.6357
1.99 | -34.1583 | -3.6469
1.9961 | -34.1584 | -3.6592
2.0022 | -34.1586 | -3.6708
2.0083 | -34.1587 | -3.6822
2.0144 | -34.1589 | -3.6937
2.0205 | -34.1592 | -3.7057
2.0266 | -34.1595 | -3.7172
2.0327 | -34.1598 | -3.729
2.0388 | -34.1602 | -3.7407
2.0449 | -34.1603 | -3.7523
2.051 | -34.1606 | -3.7638
2.0571 | -34.1609 | -3.7753
2.0632 | -34.1611 | -3.7869
2.0693 | -34.1613 | -3.7993
2.0754 | -34.1616 | -3.8108
2.0816 | -34.1619 | -3.8231
2.0877 | -34.1621 | -3.8348
2.0938 | -34.1624 | -3.8471
2.0999 | -34.1626 | -3.8589
2.106 | -34.1628 | -3.8707
2.1121 | -34.163 | -3.8828
2.1182 | -34.1632 | -3.8938
2.1243 | -34.1635 | -3.9047
2.1304 | -34.1637 | -3.916
2.1365 | -34.164 | -3.9279
2.1426 | -34.1642 | -3.94
2.1487 | -34.1645 | -3.9514
2.1548 | -34.1649 | -3.9634
2.1609 | -34.1654 | -3.975
2.167 | -34.1658 | -3.9863
2.1731 | -34.1661 | -3.9985
2.1792 | -34.1663 | -4.0097
2.1853 | -34.1667 | -4.0218
2.1914 | -34.167 | -4.0333
2.1975 | -34.1673 | -4.045
2.2036 | -34.1676 | -4.0573
2.2097 | -34.1678 | -4.0689
2.2158 | -34.1682 | -4.0797
2.222 | -34.1685 | -4.0903
2.2281 | -34.1687 | -4.1019
2.2342 | -34.169 | -4.1134
2.2403 | -34.1692 | -4.1242
2.2464 | -34.1695 | -4.1351
2.2525 | -34.1697 | -4.1466
2.2586 | -34.1701 | -4.1573
2.2647 | -34.1704 | -4.1679
2.2708 | -34.1707 | -4.1793
2.2769 | -34.1711 | -4.1901
2.283 | -34.1712 | -4.2014
2.2891 | -34.1715 | -4.2125
2.2952 | -34.1716 | -4.2231
2.3013 | -34.1719 | -4.2342
2.3074 | -34.1721 | -4.2448
2.3135 | -34.1724 | -4.2563
2.3196 | -34.1726 | -4.2666
2.3257 | -34.1728 | -4.2775
2.3318 | -34.1729 | -4.2887
2.3379 | -34.1731 | -4.2996
2.344 | -34.1733 | -4.3112
2.3501 | -34.1735 | -4.322
2.3562 | -34.1736 | -4.3325
2.3623 | -34.1738 | -4.3431
2.3685 | -34.1739 | -4.3537
2.3746 | -34.174 | -4.3651
2.3807 | -34.1742 | -4.3757
2.3868 | -34.1744 | -4.3866
2.3929 | -34.1746 | -4.3974
2.399 | -34.1747 | -4.4081
2.4051 | -34.175 | -4.4187
2.4112 | -34.1751 | -4.4295
2.4173 | -34.1753 | -4.4403
2.4234 | -34.1754 | -4.4512
2.4295 | -34.1755 | -4.4622
2.4356 | -34.1757 | -4.4732
2.4417 | -34.1758 | -4.4845
2.4478 | -34.1759 | -4.4955
2.4539 | -34.1761 | -4.5059
2.46 | -34.1762 | -4.5174
2.4661 | -34.1764 | -4.5287
2.4722 | -34.1765 | -4.539
2.4783 | -34.1766 | -4.5495
2.4844 | -34.1767 | -4.5607
2.4905 | -34.1768 | -4.5718
2.4966 | -34.1769 | -4.5827
2.5027 | -34.177 | -4.5934
2.5089 | -34.1772 | -4.6044
2.515 | -34.1772 | -4.6153
2.5211 | -34.1773 | -4.6262
2.5272 | -34.1773 | -4.6372
2.5333 | -34.1775 | -4.6484
2.5394 | -34.1776 | -4.6593
2.5455 | -34.1777 | -4.6708
2.5516 | -34.1779 | -4.6812
2.5577 | -34.1781 | -4.6928
2.5638 | -34.1782 | -4.7047
2.5699 | -34.1782 | -4.7162
2.576 | -34.1783 | -4.7279
2.5821 | -34.1784 | -4.7393
2.5882 | -34.1785 | -4.7505
2.5943 | -34.1786 | -4.7615
2.6004 | -34.1788 | -4.7726
2.6065 | -34.1789 | -4.7831
2.6126 | -34.179 | -4.7937
2.6187 | -34.1792 | -4.8043
2.6248 | -34.1793 | -4.8152
2.6309 | -34.1795 | -4.8267
2.637 | -34.1797 | -4.8385
2.6431 | -34.1798 | -4.8491
2.6492 | -34.1799 | -4.8605
2.6554 | -34.1801 | -4.8715
2.6615 | -34.1802 | -4.8826
2.6676 | -34.1805 | -4.8942
2.6737 | -34.1807 | -4.9054
2.6798 | -34.181 | -4.9166
2.6859 | -34.1812 | -4.928
2.692 | -34.1813 | -4.939
2.6981 | -34.1815 | -4.9503
2.7042 | -34.1817 | -4.9616
2.7103 | -34.1819 | -4.9723
2.7164 | -34.1822 | -4.9834
2.7225 | -34.1825 | -4.9949
2.7286 | -34.1827 | -5.006
2.7347 | -34.1829 | -5.0166
2.7408 | -34.1832 | -5.0268
2.7469 | -34.1834 | -5.0383
2.753 | -34.1836 | -5.049
2.7591 | -34.1839 | -5.0604
2.7652 | -34.1842 | -5.072
2.7713 | -34.1843 | -5.0828
2.7774 | -34.1846 | -5.0939
2.7835 | -34.1847 | -5.105
2.7896 | -34.1849 | -5.1156
2.7958 | -34.1852 | -5.1265
2.8019 | -34.1855 | -5.137
2.808 | -34.1858 | -5.1479
2.8141 | -34.1861 | -5.1592
2.8202 | -34.1864 | -5.1702
2.8263 | -34.1867 | -5.1812
2.8324 | -34.187 | -5.1927
2.8385 | -34.1874 | -5.2036
2.8446 | -34.1878 | -5.2145
2.8507 | -34.1882 | -5.2255
2.8568 | -34.1885 | -5.236
2.8629 | -34.1888 | -5.2468
2.869 | -34.1891 | -5.258
2.8751 | -34.1894 | -5.2689
2.8812 | -34.1898 | -5.2803
2.8873 | -34.1901 | -5.2914
2.8934 | -34.1904 | -5.3025
2.8995 | -34.1908 | -5.3138
2.9056 | -34.1911 | -5.3251
2.9117 | -34.1914 | -5.3362
2.9178 | -34.1918 | -5.3471
2.9239 | -34.1922 | -5.3588
2.93 | -34.1926 | -5.3704
2.9361 | -34.1929 | -5.3813
2.9423 | -34.1932 | -5.3921
2.9484 | -34.1936 | -5.4033
2.9545 | -34.194 | -5.4145
2.9606 | -34.1944 | -5.426
2.9667 | -34.1948 | -5.4374
2.9728 | -34.1951 | -5.4488
2.9789 | -34.1956 | -5.46
2.985 | -34.196 | -5.4715
2.9911 | -34.1964 | -5.483
2.9972 | -34.1968 | -5.4942
3.0033 | -34.1972 | -5.5059
3.0094 | -34.1976 | -5.5171
3.0155 | -34.1981 | -5.5282
3.0216 | -34.1984 | -5.5397
3.0277 | -34.1987 | -5.5515
3.0338 | -34.1992 | -5.5631
3.0399 | -34.1996 | -5.5745
3.046 | -34.1999 | -5.5863
3.0521 | -34.2002 | -5.5976
3.0582 | -34.2005 | -5.6099
3.0643 | -34.2009 | -5.6215
3.0704 | -34.2012 | -5.6335
3.0765 | -34.2015 | -5.6448
3.0827 | -34.2018 | -5.6562
3.0888 | -34.2022 | -5.6677
3.0949 | -34.2026 | -5.6792
3.101 | -34.2029 | -5.6909
3.1071 | -34.2032 | -5.7028
3.1132 | -34.2036 | -5.7148
3.1193 | -34.204 | -5.7264
3.1254 | -34.2044 | -5.738
3.1315 | -34.2047 | -5.7501
3.1376 | -34.205 | -5.7617
3.1437 | -34.2054 | -5.7735
3.1498 | -34.2057 | -5.7857
3.1559 | -34.206 | -5.7973
3.162 | -34.2063 | -5.8097
3.1681 | -34.2066 | -5.8218
3.1742 | -34.2069 | -5.8336
3.1803 | -34.2072 | -5.8454
3.1864 | -34.2076 | -5.8569
3.1925 | -34.2079 | -5.8687
3.1986 | -34.2082 | -5.8809
3.2047 | -34.2085 | -5.8931
3.2108 | -34.2087 | -5.9044
3.2169 | -34.2091 | -5.9164
3.223 | -34.2093 | -5.9283
3.2292 | -34.2095 | -5.9403
3.2353 | -34.2097 | -5.9514
3.2414 | -34.21 | -5.963
3.2475 | -34.2102 | -5.9746
3.2536 | -34.2105 | -5.9863
3.2597 | -34.2108 | -5.9974
3.2658 | -34.211 | -6.0085
3.2719 | -34.2112 | -6.0198
3.278 | -34.2115 | -6.0316
3.2841 | -34.2117 | -6.0429
3.2902 | -34.212 | -6.0545
3.2963 | -34.2122 | -6.0662
3.3024 | -34.2123 | -6.0775
3.3085 | -34.2126 | -6.0892
3.3146 | -34.2128 | -6.1006
3.3207 | -34.2131 | -6.1123
3.3268 | -34.2133 | -6.1239
3.3329 | -34.2135 | -6.1349
3.339 | -34.2137 | -6.1459
3.3451 | -34.2138 | -6.1573
3.3512 | -34.214 | -6.1682
3.3573 | -34.2143 | -6.1797
3.3634 | -34.2144 | -6.191
3.3696 | -34.2147 | -6.2025
3.3757 | -34.2148 | -6.2135
3.3818 | -34.215 | -6.2246
3.3879 | -34.2152 | -6.2359
3.394 | -34.2154 | -6.2474
3.4001 | -34.2156 | -6.2584
3.4062 | -34.2157 | -6.2695
3.4123 | -34.2158 | -6.2803
3.4184 | -34.216 | -6.2918
3.4245 | -34.2162 | -6.3027
3.4306 | -34.2163 | -6.3136
3.4367 | -34.2165 | -6.3244
3.4428 | -34.2167 | -6.3349
3.4489 | -34.2168 | -6.346
3.455 | -34.2169 | -6.3571
3.4611 | -34.2171 | -6.3685
3.4672 | -34.2173 | -6.3797
3.4733 | -34.2175 | -6.391
3.4794 | -34.2176 | -6.402
3.4855 | -34.2178 | -6.4126
3.4916 | -34.218 | -6.4231
3.4977 | -34.2181 | -6.4346
3.5038 | -34.2183 | -6.4456
3.5099 | -34.2184 | -6.4561
3.5161 | -34.2186 | -6.4666
3.5222 | -34.2187 | -6.4771
3.5283 | -34.2188 | -6.4881
3.5344 | -34.2189 | -6.4991
3.5405 | -34.2191 | -6.5098
3.5466 | -34.2193 | -6.5209
3.5527 | -34.2194 | -6.5318
3.5588 | -34.2195 | -6.5426
3.5649 | -34.2195 | -6.5534
3.571 | -34.2197 | -6.5642
3.5771 | -34.2198 | -6.5755
3.5832 | -34.2199 | -6.5855
3.5893 | -34.2202 | -6.5963
3.5954 | -34.2203 | -6.6073
3.6015 | -34.2205 | -6.6184
3.6076 | -34.2207 | -6.6291
3.6137 | -34.2208 | -6.6393
3.6198 | -34.2211 | -6.6497
3.6259 | -34.2212 | -6.6608
3.632 | -34.2214 | -6.6715
3.6381 | -34.2216 | -6.6825
3.6442 | -34.2218 | -6.6932
3.6503 | -34.222 | -6.7035
3.6565 | -34.2222 | -6.7145
3.6626 | -34.2223 | -6.7252
3.6687 | -34.2225 | -6.7361
3.6748 | -34.2228 | -6.7467
3.6809 | -34.223 | -6.7575
3.687 | -34.2232 | -6.7686
3.6931 | -34.2235 | -6.779
3.6992 | -34.2237 | -6.7897
3.7053 | -34.2239 | -6.8004
3.7114 | -34.224 | -6.8112
3.7175 | -34.2243 | -6.8221
3.7236 | -34.2245 | -6.8323
3.7297 | -34.2247 | -6.8432
3.7358 | -34.2249 | -6.8543
3.7419 | -34.2251 | -6.8653
3.748 | -34.2254 | -6.876
3.7541 | -34.2256 | -6.8867
3.7602 | -34.2259 | -6.8974
3.7663 | -34.2261 | -6.908
3.7724 | -34.2263 | -6.9188
3.7785 | -34.2266 | -6.9293
3.7846 | -34.2268 | -6.9399
3.7907 | -34.2271 | -6.951
3.7969 | -34.2274 | -6.9619
3.803 | -34.2276 | -6.9723
3.8091 | -34.2279 | -6.983
3.8152 | -34.2282 | -6.9936
3.8213 | -34.2284 | -7.0046
3.8274 | -34.2287 | -7.0151
3.8335 | -34.229 | -7.026
3.8396 | -34.2293 | -7.0367
3.8457 | -34.2296 | -7.0476
3.8518 | -34.23 | -7.0583
3.8579 | -34.2303 | -7.069
3.864 | -34.2306 | -7.08
3.8701 | -34.2309 | -7.0911
3.8762 | -34.2312 | -7.1017
3.8823 | -34.2314 | -7.1127
3.8884 | -34.2318 | -7.1235
3.8945 | -34.2321 | -7.1346
3.9006 | -34.2324 | -7.1457
3.9067 | -34.2327 | -7.157
3.9128 | -34.233 | -7.1679
3.9189 | -34.2333 | -7.1784
3.925 | -34.2336 | -7.1894
3.9311 | -34.2338 | -7.1998
3.9372 | -34.2342 | -7.2108
3.9434 | -34.2345 | -7.2217
3.9495 | -34.2349 | -7.2326
3.9556 | -34.2352 | -7.2433
3.9617 | -34.2355 | -7.2537
3.9678 | -34.2359 | -7.2643
3.9739 | -34.2361 | -7.2751
3.98 | -34.2364 | -7.286
3.9861 | -34.2367 | -7.2968
3.9922 | -34.2369 | -7.3076
3.9983 | -34.2373 | -7.318
4.0044 | -34.2376 | -7.3284
4.0105 | -34.2379 | -7.3393
4.0166 | -34.2382 | -7.3501
4.0227 | -34.2385 | -7.3609
4.0288 | -34.2388 | -7.3722
4.0349 | -34.239 | -7.3828
4.041 | -34.2393 | -7.3935
4.0471 | -34.2395 | -7.4044
4.0532 | -34.2399 | -7.4149
4.0593 | -34.2402 | -7.4259
4.0654 | -34.2405 | -7.4364
4.0715 | -34.2408 | -7.4477
4.0776 | -34.2411 | -7.4584
4.0838 | -34.2414 | -7.4696
4.0899 | -34.2417 | -7.4801
4.096 | -34.242 | -7.4908
4.1021 | -34.2423 | -7.5014
4.1082 | -34.2426 | -7.5119
4.1143 | -34.2428 | -7.5224
4.1204 | -34.2432 | -7.5328
4.1265 | -34.2435 | -7.5433
4.1326 | -34.2437 | -7.5537
4.1387 | -34.2441 | -7.5643
4.1448 | -34.2444 | -7.5751
4.1509 | -34.2446 | -7.5858
4.157 | -34.2449 | -7.5969
4.1631 | -34.2451 | -7.6079
4.1692 | -34.2454 | -7.6184
4.1753 | -34.2457 | -7.6291
4.1814 | -34.2459 | -7.6394
4.1875 | -34.2463 | -7.6504
4.1936 | -34.2465 | -7.6612
4.1997 | -34.2468 | -7.6716
4.2058 | -34.247 | -7.6825
4.2119 | -34.2473 | -7.693
4.218 | -34.2475 | -7.7034
4.2241 | -34.2478 | -7.7139
4.2303 | -34.2481 | -7.7244
4.2364 | -34.2483 | -7.735
4.2425 | -34.2485 | -7.7456
4.2486 | -34.2488 | -7.7564
4.2547 | -34.2489 | -7.7668
4.2608 | -34.2492 | -7.7776
4.2669 | -34.2494 | -7.7879
4.273 | -34.2496 | -7.7986
4.2791 | -34.2499 | -7.8093
4.2852 | -34.2501 | -7.82
4.2913 | -34.2503 | -7.8303
4.2974 | -34.2505 | -7.8412
4.3035 | -34.2508 | -7.8521
4.3096 | -34.251 | -7.8629
4.3157 | -34.2513 | -7.8733
4.3218 | -34.2515 | -7.8838
4.3279 | -34.2517 | -7.8944
4.334 | -34.2519 | -7.9049
4.3401 | -34.2521 | -7.9151
4.3462 | -34.2524 | -7.9257
4.3523 | -34.2527 | -7.9359
4.3584 | -34.2529 | -7.946
4.3645 | -34.2531 | -7.9565
4.3707 | -34.2533 | -7.9667
4.3768 | -34.2535 | -7.9772
4.3829 | -34.2538 | -7.9881
4.389 | -34.254 | -7.9986
4.3951 | -34.2542 | -8.0089
4.4012 | -34.2544 | -8.0197
4.4073 | -34.2546 | -8.0301
4.4134 | -34.2548 | -8.0405
4.4195 | -34.2551 | -8.0509
4.4256 | -34.2554 | -8.0612
4.4317 | -34.2556 | -8.0716
4.4378 | -34.2558 | -8.0817
4.4439 | -34.256 | -8.0919
4.45 | -34.2562 | -8.1023
4.4561 | -34.2563 | -8.1128
4.4622 | -34.2565 | -8.1233
4.4683 | -34.2568 | -8.1337
4.4744 | -34.257 | -8.1445
4.4805 | -34.2572 | -8.1545
4.4866 | -34.2574 | -8.1649
4.4927 | -34.2577 | -8.1756
4.4988 | -34.2579 | -8.186
4.5049 | -34.2581 | -8.196
4.511 | -34.2583 | -8.2062
4.5172 | -34.2586 | -8.2165
4.5233 | -34.2587 | -8.2273
4.5294 | -34.259 | -8.2376
4.5355 | -34.2592 | -8.248
4.5416 | -34.2594 | -8.2583
4.5477 | -34.2596 | -8.2687
4.5538 | -34.2598 | -8.2787
4.5599 | -34.2601 | -8.289
4.566 | -34.2603 | -8.2994
4.5721 | -34.2605 | -8.3098
4.5782 | -34.2608 | -8.3205
4.5843 | -34.261 | -8.3309
4.5904 | -34.2613 | -8.3414
4.5965 | -34.2615 | -8.3516
4.6026 | -34.2617 | -8.3625
4.6087 | -34.2619 | -8.3728
4.6148 | -34.2621 | -8.3835
4.6209 | -34.2624 | -8.3938
4.627 | -34.2626 | -8.4044
4.6331 | -34.2628 | -8.4147
4.6392 | -34.263 | -8.4252
4.6453 | -34.2632 | -8.4356
4.6514 | -34.2635 | -8.446
4.6576 | -34.2637 | -8.4568
4.6637 | -34.2639 | -8.4668
4.6698 | -34.2641 | -8.4771
4.6759 | -34.2643 | -8.4879
4.682 | -34.2645 | -8.4981
4.6881 | -34.2646 | -8.5085
4.6942 | -34.2649 | -8.5189
4.7003 | -34.2651 | -8.5294
4.7064 | -34.2653 | -8.5395
4.7125 | -34.2656 | -8.5502
4.7186 | -34.2658 | -8.5604
4.7247 | -34.2661 | -8.5709
4.7308 | -34.2663 | -8.5812
4.7369 | -34.2665 | -8.5917
4.743 | -34.2667 | -8.6025
4.7491 | -34.267 | -8.6132
4.7552 | -34.2672 | -8.6238
4.7613 | -34.2674 | -8.6343
4.7674 | -34.2676 | -8.6446
4.7735 | -34.2678 | -8.6555
4.7796 | -34.2681 | -8.6662
4.7857 | -34.2683 | -8.6768
4.7918 | -34.2685 | -8.6871
4.7979 | -34.2688 | -8.6974
4.8041 | -34.269 | -8.7079
4.8102 | -34.2692 | -8.718
4.8163 | -34.2695 | -8.7284
4.8224 | -34.2697 | -8.7389
4.8285 | -34.27 | -8.7491
4.8346 | -34.2702 | -8.7593
4.8407 | -34.2704 | -8.7698
4.8468 | -34.2707 | -8.7806
4.8529 | -34.2709 | -8.7913
4.859 | -34.2712 | -8.8021
4.8651 | -34.2714 | -8.8123
4.8712 | -34.2717 | -8.8226
4.8773 | -34.2718 | -8.8331
4.8834 | -34.2721 | -8.8436
4.8895 | -34.2723 | -8.8539
4.8956 | -34.2725 | -8.8643
4.9017 | -34.2727 | -8.8746
4.9078 | -34.273 | -8.8854
4.9139 | -34.2732 | -8.8954
4.92 | -34.2735 | -8.9056
4.9261 | -34.2738 | -8.9157
4.9322 | -34.2741 | -8.9262
4.9383 | -34.2743 | -8.9367
4.9445 | -34.2746 | -8.9471
4.9506 | -34.2748 | -8.9574
4.9567 | -34.275 | -8.9676
4.9628 | -34.2753 | -8.978
4.9689 | -34.2755 | -8.9888
4.975 | -34.2757 | -8.9996
4.9811 | -34.276 | -9.0098
4.9872 | -34.2763 | -9.0201
4.9933 | -34.2766 | -9.0308
4.9994 | -34.2769 | -9.0417
5.0055 | -34.2772 | -9.0522
5.0116 | -34.2774 | -9.0631
5.0177 | -34.2777 | -9.0739
5.0238 | -34.278 | -9.0846
5.0299 | -34.2784 | -9.0955
5.036 | -34.2786 | -9.106
5.0421 | -34.2788 | -9.1164
5.0482 | -34.2791 | -9.1269
5.0543 | -34.2794 | -9.1371
5.0604 | -34.2797 | -9.1476
5.0665 | -34.28 | -9.1579
5.0726 | -34.2803 | -9.1687
5.0787 | -34.2806 | -9.1794
5.0848 | -34.2808 | -9.19
5.091 | -34.2811 | -9.2003
5.0971 | -34.2814 | -9.2105
5.1032 | -34.2817 | -9.2213
5.1093 | -34.282 | -9.2319
5.1154 | -34.2822 | -9.2421
5.1215 | -34.2825 | -9.2524
5.1276 | -34.2827 | -9.2627
5.1337 | -34.2831 | -9.2729
5.1398 | -34.2833 | -9.2831
5.1459 | -34.2836 | -9.2939
5.152 | -34.2839 | -9.3046
5.1581 | -34.2842 | -9.3153
5.1642 | -34.2845 | -9.3259
5.1703 | -34.2848 | -9.3364
5.1764 | -34.2851 | -9.3467
5.1825 | -34.2854 | -9.357
5.1886 | -34.2857 | -9.3672
5.1947 | -34.2859 | -9.3777
5.2008 | -34.2863 | -9.3884
5.2069 | -34.2865 | -9.3991
5.213 | -34.2868 | -9.4094
5.2191 | -34.2871 | -9.42
5.2252 | -34.2874 | -9.4305
5.2314 | -34.2877 | -9.4412
5.2375 | -34.288 | -9.4515
5.2436 | -34.2882 | -9.4621
5.2497 | -34.2886 | -9.4723
5.2558 | -34.2889 | -9.4832
5.2619 | -34.2892 | -9.494
5.268 | -34.2895 | -9.5047
5.2741 | -34.2898 | -9.5153
5.2802 | -34.2901 | -9.5262
5.2863 | -34.2904 | -9.5367
5.2924 | -34.2907 | -9.5474
5.2985 | -34.2911 | -9.558
5.3046 | -34.2914 | -9.5686
5.3107 | -34.2917 | -9.5794
5.3168 | -34.292 | -9.59
5.3229 | -34.2923 | -9.6006
5.329 | -34.2926 | -9.6112
5.3351 | -34.2929 | -9.6215
5.3412 | -34.2932 | -9.632
5.3473 | -34.2935 | -9.6426
5.3534 | -34.2938 | -9.6533
5.3595 | -34.2941 | -9.664
5.3656 | -34.2945 | -9.6745
5.3717 | -34.2947 | -9.6855
5.3779 | -34.2951 | -9.6964
5.384 | -34.2954 | -9.707
5.3901 | -34.2957 | -9.7178
5.3962 | -34.296 | -9.7282
5.4023 | -34.2963 | -9.7391
5.4084 | -34.2967 | -9.7496
5.4145 | -34.297 | -9.76
5.4206 | -34.2973 | -9.7706
5.4267 | -34.2976 | -9.7815
5.4328 | -34.2979 | -9.7922
5.4389 | -34.2982 | -9.8025
5.445 | -34.2985 | -9.8135
5.4511 | -34.2989 | -9.8239
5.4572 | -34.2992 | -9.8346
5.4633 | -34.2995 | -9.8452
5.4694 | -34.2998 | -9.8559
5.4755 | -34.3001 | -9.8662
5.4816 | -34.3004 | -9.8769
5.4877 | -34.3008 | -9.8874
5.4938 | -34.3011 | -9.8977
5.4999 | -34.3014 | -9.9082
5.506 | -34.3018 | -9.9188
5.5121 | -34.302 | -9.9293
5.5183 | -34.3024 | -9.9397
5.5244 | -34.3027 | -9.9506
5.5305 | -34.303 | -9.9611
5.5366 | -34.3033 | -9.9715
5.5427 | -34.3036 | -9.9822
5.5488 | -34.3039 | -9.9929
5.5549 | -34.3042 | -10.0036
5.561 | -34.3045 | -10.0144
5.5671 | -34.3049 | -10.0249
5.5732 | -34.3052 | -10.0357
5.5793 | -34.3055 | -10.0467
5.5854 | -34.3059 | -10.0573
5.5915 | -34.3062 | -10.0682
5.5976 | -34.3065 | -10.0789
5.6037 | -34.3068 | -10.0896
5.6098 | -34.3071 | -10.1002
5.6159 | -34.3074 | -10.1106
5.622 | -34.3077 | -10.1214
5.6281 | -34.3081 | -10.1321
5.6342 | -34.3084 | -10.143
5.6403 | -34.3087 | -10.1537
5.6464 | -34.309 | -10.1645
5.6525 | -34.3093 | -10.1752
5.6586 | -34.3097 | -10.1856
5.6648 | -34.31 | -10.1963
5.6709 | -34.3104 | -10.2071
5.677 | -34.3107 | -10.2178
5.6831 | -34.311 | -10.2285
5.6892 | -34.3113 | -10.2393
5.6953 | -34.3116 | -10.25
5.7014 | -34.312 | -10.2608
5.7075 | -34.3122 | -10.2714
5.7136 | -34.3126 | -10.282
5.7197 | -34.3129 | -10.2926
5.7258 | -34.3132 | -10.3031
5.7319 | -34.3135 | -10.3138
5.738 | -34.3139 | -10.3244
5.7441 | -34.3142 | -10.335
5.7502 | -34.3146 | -10.3454
5.7563 | -34.3149 | -10.356
5.7624 | -34.3152 | -10.3669
5.7685 | -34.3156 | -10.3776
5.7746 | -34.3158 | -10.3885
5.7807 | -34.3161 | -10.3992
5.7868 | -34.3164 | -10.4101
5.7929 | -34.3167 | -10.4206
5.799 | -34.317 | -10.4312
5.8052 | -34.3174 | -10.4417
5.8113 | -34.3177 | -10.4522
5.8174 | -34.318 | -10.4627
5.8235 | -34.3183 | -10.4733
5.8296 | -34.3186 | -10.4842
5.8357 | -34.3189 | -10.4948
5.8418 | -34.3192 | -10.5055
5.8479 | -34.3195 | -10.5163
5.854 | -34.3198 | -10.527
5.8601 | -34.3202 | -10.5375
5.8662 | -34.3205 | -10.5483
5.8723 | -34.3208 | -10.5588
5.8784 | -34.3211 | -10.5693
5.8845 | -34.3214 | -10.5799
5.8906 | -34.3217 | -10.5905
5.8967 | -34.322 | -10.601
5.9028 | -34.3224 | -10.6114
5.9089 | -34.3227 | -10.6219
5.915 | -34.323 | -10.6323
5.9211 | -34.3233 | -10.6429
5.9272 | -34.3236 | -10.6532
5.9333 | -34.324 | -10.6637
5.9394 | -34.3243 | -10.6743
5.9455 | -34.3246 | -10.6847
5.9517 | -34.325 | -10.6951
5.9578 | -34.3254 | -10.7058
5.9639 | -34.3257 | -10.7161
5.97 | -34.326 | -10.7267
5.9761 | -34.3264 | -10.7372
5.9822 | -34.3267 | -10.7478
5.9883 | -34.327 | -10.7584
5.9944 | -34.3273 | -10.7691
6.0005 | -34.3277 | -10.7798
6.0066 | -34.328 | -10.7902
6.0127 | -34.3284 | -10.8005
6.0188 | -34.3287 | -10.8113
6.0249 | -34.329 | -10.822
6.031 | -34.3293 | -10.8324
6.0371 | -34.3296 | -10.8429
6.0432 | -34.3299 | -10.8534
6.0493 | -34.3303 | -10.8637
6.0554 | -34.3306 | -10.8741
6.0615 | -34.331 | -10.8846
6.0676 | -34.3313 | -10.8954
6.0737 | -34.3317 | -10.906
6.0798 | -34.3319 | -10.9165
6.0859 | -34.3323 | -10.927
6.0921 | -34.3326 | -10.9373
6.0982 | -34.333 | -10.9478
6.1043 | -34.3333 | -10.9582
6.1104 | -34.3337 | -10.9689
6.1165 | -34.334 | -10.9789
6.1226 | -34.3344 | -10.9894
6.1287 | -34.3347 | -10.9998
6.1348 | -34.335 | -11.0103
6.1409 | -34.3354 | -11.0207
6.147 | -34.3357 | -11.0311
6.1531 | -34.336 | -11.0415
6.1592 | -34.3363 | -11.0519
6.1653 | -34.3366 | -11.0622
6.1714 | -34.3369 | -11.0727
6.1775 | -34.3373 | -11.0833
6.1836 | -34.3376 | -11.0938
6.1897 | -34.338 | -11.1042
6.1958 | -34.3383 | -11.1149
6.2019 | -34.3386 | -11.1255
6.208 | -34.339 | -11.136
6.2141 | -34.3393 | -11.1464
6.2202 | -34.3397 | -11.1569
6.2263 | -34.34 | -11.1673
6.2325 | -34.3403 | -11.1774
6.2386 | -34.3407 | -11.1875
6.2447 | -34.341 | -11.1982
6.2508 | -34.3414 | -11.2086
6.2569 | -34.3417 | -11.219
6.263 | -34.342 | -11.2294
6.2691 | -34.3424 | -11.2398
6.2752 | -34.3427 | -11.2503
6.2813 | -34.343 | -11.2606
6.2874 | -34.3433 | -11.2709
6.2935 | -34.3436 | -11.2813
6.2996 | -34.3439 | -11.2917
6.3057 | -34.3443 | -11.3021
6.3118 | -34.3446 | -11.3126
6.3179 | -34.3449 | -11.323
6.324 | -34.3452 | -11.3334
6.3301 | -34.3456 | -11.3437
6.3362 | -34.3459 | -11.354
6.3423 | -34.3463 | -11.3645
6.3484 | -34.3466 | -11.375
6.3545 | -34.3469 | -11.385
6.3606 | -34.3472 | -11.3951
6.3667 | -34.3475 | -11.4053
6.3728 | -34.3479 | -11.4158
6.379 | -34.3482 | -11.4261
6.3851 | -34.3486 | -11.4363
6.3912 | -34.3488 | -11.4465
6.3973 | -34.3492 | -11.4569
6.4034 | -34.3495 | -11.4671
6.4095 | -34.3498 | -11.4774
6.4156 | -34.3501 | -11.4878
6.4217 | -34.3504 | -11.498
6.4278 | -34.3508 | -11.5085
6.4339 | -34.3511 | -11.5187
6.44 | -34.3515 | -11.529
6.4461 | -34.3518 | -11.5395
6.4522 | -34.3521 | -11.5496
6.4583 | -34.3524 | -11.5601
6.4644 | -34.3527 | -11.5705
6.4705 | -34.353 | -11.5805
6.4766 | -34.3534 | -11.5908
6.4827 | -34.3537 | -11.6012
6.4888 | -34.354 | -11.6113
6.4949 | -34.3543 | -11.6213
6.501 | -34.3547 | -11.6317
6.5071 | -34.355 | -11.6418
6.5132 | -34.3553 | -11.6521
6.5194 | -34.3556 | -11.6624
6.5255 | -34.356 | -11.6725
6.5316 | -34.3563 | -11.6829
6.5377 | -34.3566 | -11.6931
6.5438 | -34.3569 | -11.7032
6.5499 | -34.3572 | -11.7133
6.556 | -34.3575 | -11.7236
6.5621 | -34.3578 | -11.7341
6.5682 | -34.3582 | -11.7444
6.5743 | -34.3586 | -11.7546
6.5804 | -34.3589 | -11.7649
6.5865 | -34.3592 | -11.7752
6.5926 | -34.3595 | -11.7854
6.5987 | -34.3599 | -11.7953
6.6048 | -34.3602 | -11.8056
6.6109 | -34.3605 | -11.8158
6.617 | -34.3608 | -11.8262
6.6231 | -34.3611 | -11.8364
6.6292 | -34.3615 | -11.8466
6.6353 | -34.3618 | -11.8572
6.6414 | -34.3621 | -11.8673
6.6475 | -34.3625 | -11.8775
6.6536 | -34.3628 | -11.8876
6.6597 | -34.3631 | -11.8976
6.6659 | -34.3634 | -11.9078
6.672 | -34.3638 | -11.9179
6.6781 | -34.3641 | -11.9282
6.6842 | -34.3644 | -11.9382
6.6903 | -34.3648 | -11.9484
6.6964 | -34.3651 | -11.9586
6.7025 | -34.3654 | -11.9686
6.7086 | -34.3657 | -11.9788
6.7147 | -34.3661 | -11.9891
6.7208 | -34.3664 | -11.9994
6.7269 | -34.3667 | -12.0095
6.733 | -34.367 | -12.0195
6.7391 | -34.3673 | -12.0297
6.7452 | -34.3677 | -12.0396
6.7513 | -34.368 | -12.0495
6.7574 | -34.3683 | -12.0598
6.7635 | -34.3687 | -12.0697
6.7696 | -34.369 | -12.0798
6.7757 | -34.3694 | -12.09
6.7818 | -34.3697 | -12.1001
6.7879 | -34.3701 | -12.1101
6.794 | -34.3704 | -12.1201
6.8001 | -34.3707 | -12.1301
6.8063 | -34.371 | -12.1403
6.8124 | -34.3713 | -12.1505
6.8185 | -34.3716 | -12.1609
6.8246 | -34.3719 | -12.1708
6.8307 | -34.3722 | -12.1808
6.8368 | -34.3726 | -12.1912
6.8429 | -34.3729 | -12.2013
6.849 | -34.3733 | -12.2115
6.8551 | -34.3736 | -12.2215
6.8612 | -34.374 | -12.2316
6.8673 | -34.3743 | -12.2417
6.8734 | -34.3746 | -12.2519
6.8795 | -34.3749 | -12.2621
6.8856 | -34.3753 | -12.2722
6.8917 | -34.3756 | -12.2822
6.8978 | -34.3759 | -12.292
6.9039 | -34.3762 | -12.3019
6.91 | -34.3766 | -12.312
6.9161 | -34.3769 | -12.3219
6.9222 | -34.3772 | -12.332
6.9283 | -34.3776 | -12.3418
6.9344 | -34.378 | -12.3519
6.9405 | -34.3783 | -12.3619
6.9466 | -34.3787 | -12.3722
6.9528 | -34.379 | -12.3825
6.9589 | -34.3794 | -12.3927
6.965 | -34.3797 | -12.4028
6.9711 | -34.3801 | -12.413
6.9772 | -34.3805 | -12.4231
6.9833 | -34.3808 | -12.4331
6.9894 | -34.3811 | -12.4434
6.9955 | -34.3815 | -12.4533
7.0016 | -34.3818 | -12.4633
7.0077 | -34.3822 | -12.4733
7.0138 | -34.3825 | -12.4833
7.0199 | -34.3828 | -12.4933
7.026 | -34.3832 | -12.5033
7.0321 | -34.3835 | -12.5131
7.0382 | -34.3839 | -12.5229
7.0443 | -34.3842 | -12.5331
7.0504 | -34.3846 | -12.5432
7.0565 | -34.3849 | -12.5533
7.0626 | -34.3853 | -12.5631
7.0687 | -34.3856 | -12.5727
7.0748 | -34.386 | -12.5825
7.0809 | -34.3863 | -12.5925
7.087 | -34.3867 | -12.6026
7.0932 | -34.387 | -12.6126
7.0993 | -34.3873 | -12.6226
7.1054 | -34.3877 | -12.6325
7.1115 | -34.3881 | -12.6422
7.1176 | -34.3884 | -12.6522
7.1237 | -34.3887 | -12.6621
7.1298 | -34.3891 | -12.672
7.1359 | -34.3894 | -12.682
7.142 | -34.3898 | -12.6918
7.1481 | -34.3901 | -12.7016
7.1542 | -34.3904 | -12.7117
7.1603 | -34.3908 | -12.7216
7.1664 | -34.3912 | -12.7315
7.1725 | -34.3915 | -12.7416
7.1786 | -34.3919 | -12.7515
7.1847 | -34.3922 | -12.7616
7.1908 | -34.3926 | -12.7717
7.1969 | -34.3929 | -12.7819
7.203 | -34.3933 | -12.7916
7.2091 | -34.3936 | -12.8015
7.2152 | -34.394 | -12.8115
7.2213 | -34.3943 | -12.8213
7.2274 | -34.3947 | -12.8312
7.2335 | -34.395 | -12.8411
7.2397 | -34.3954 | -12.851
7.2458 | -34.3958 | -12.8608
7.2519 | -34.3961 | -12.8707
7.258 | -34.3965 | -12.8808
7.2641 | -34.3968 | -12.8908
7.2702 | -34.3972 | -12.9009
7.2763 | -34.3976 | -12.9109
7.2824 | -34.3979 | -12.921
7.2885 | -34.3983 | -12.931
7.2946 | -34.3986 | -12.941
7.3007 | -34.3989 | -12.951
7.3068 | -34.3993 | -12.9612
7.3129 | -34.3997 | -12.9713
7.319 | -34.4 | -12.9813
7.3251 | -34.4004 | -12.9911
7.3312 | -34.4007 | -13.001
7.3373 | -34.4011 | -13.0108
7.3434 | -34.4015 | -13.0206
7.3495 | -34.4019 | -13.0305
7.3556 | -34.4022 | -13.0405
7.3617 | -34.4026 | -13.0506
7.3678 | -34.4029 | -13.0605
7.3739 | -34.4033 | -13.0707
7.3801 | -34.4036 | -13.0809
7.3862 | -34.404 | -13.0908
7.3923 | -34.4044 | -13.1009
7.3984 | -34.4047 | -13.1106
7.4045 | -34.4051 | -13.1205
7.4106 | -34.4055 | -13.1304
7.4167 | -34.4059 | -13.1405
7.4228 | -34.4063 | -13.1505
7.4289 | -34.4066 | -13.1605
7.435 | -34.407 | -13.1702
7.4411 | -34.4074 | -13.1803
7.4472 | -34.4077 | -13.1903
7.4533 | -34.4081 | -13.2001
7.4594 | -34.4085 | -13.2103
7.4655 | -34.4088 | -13.22
7.4716 | -34.4092 | -13.2301
7.4777 | -34.4095 | -13.2401
7.4838 | -34.4099 | -13.2499
7.4899 | -34.4103 | -13.2598
7.496 | -34.4106 | -13.2696
7.5021 | -34.411 | -13.2795
7.5082 | -34.4114 | -13.2894
7.5143 | -34.4117 | -13.2994
7.5204 | -34.4121 | -13.3095
7.5266 | -34.4125 | -13.3193
7.5327 | -34.4128 | -13.3292
7.5388 | -34.4132 | -13.3395
7.5449 | -34.4136 | -13.3496
7.551 | -34.414 | -13.3594
7.5571 | -34.4143 | -13.3695
7.5632 | -34.4147 | -13.3795
7.5693 | -34.4151 | -13.3894
7.5754 | -34.4155 | -13.3992
7.5815 | -34.4158 | -13.4092
7.5876 | -34.4162 | -13.4194
7.5937 | -34.4166 | -13.4295
7.5998 | -34.417 | -13.4394
7.6059 | -34.4174 | -13.4495
7.612 | -34.4178 | -13.4595
7.6181 | -34.4181 | -13.4694
7.6242 | -34.4185 | -13.4796
7.6303 | -34.4189 | -13.4898
7.6364 | -34.4192 | -13.4999
7.6425 | -34.4196 | -13.5098
7.6486 | -34.42 | -13.5199
7.6547 | -34.4204 | -13.5297
7.6608 | -34.4207 | -13.5397
7.667 | -34.4211 | -13.5498
7.6731 | -34.4215 | -13.5598
7.6792 | -34.4219 | -13.5701
7.6853 | -34.4222 | -13.5801
7.6914 | -34.4226 | -13.5903
7.6975 | -34.423 | -13.6003
7.7036 | -34.4234 | -13.6104
7.7097 | -34.4238 | -13.6206
7.7158 | -34.4242 | -13.6307
7.7219 | -34.4246 | -13.6406
7.728 | -34.4249 | -13.6507
7.7341 | -34.4253 | -13.6609
7.7402 | -34.4257 | -13.6709
7.7463 | -34.4261 | -13.681
7.7524 | -34.4264 | -13.6913
7.7585 | -34.4268 | -13.7014
7.7646 | -34.4272 | -13.7114
7.7707 | -34.4275 | -13.7216
7.7768 | -34.4279 | -13.7318
7.7829 | -34.4283 | -13.7421
7.789 | -34.4286 | -13.7524
7.7951 | -34.429 | -13.7625
7.8012 | -34.4294 | -13.7725
7.8073 | -34.4298 | -13.7828
7.8135 | -34.4302 | -13.7929
7.8196 | -34.4306 | -13.8032
7.8257 | -34.431 | -13.8135
7.8318 | -34.4313 | -13.8236
7.8379 | -34.4317 | -13.8336
7.844 | -34.4321 | -13.8436
7.8501 | -34.4325 | -13.8538
7.8562 | -34.4329 | -13.8641
7.8623 | -34.4333 | -13.8741
7.8684 | -34.4336 | -13.8843
7.8745 | -34.434 | -13.8944
7.8806 | -34.4344 | -13.9047
7.8867 | -34.4348 | -13.9146
7.8928 | -34.4352 | -13.9244
7.8989 | -34.4356 | -13.9346
7.905 | -34.4359 | -13.9448
7.9111 | -34.4363 | -13.9549
7.9172 | -34.4367 | -13.965
7.9233 | -34.4371 | -13.9752
7.9294 | -34.4375 | -13.9854
7.9355 | -34.4378 | -13.9953
7.9416 | -34.4382 | -14.0054
7.9477 | -34.4386 | -14.0154
7.9539 | -34.4389 | -14.0255
7.96 | -34.4393 | -14.0357
7.9661 | -34.4397 | -14.0459
7.9722 | -34.4401 | -14.0559
7.9783 | -34.4405 | -14.0662
7.9844 | -34.4409 | -14.0763
7.9905 | -34.4413 | -14.0866
7.9966 | -34.4417 | -14.0969
8.0027 | -34.4421 | -14.107
8.0088 | -34.4424 | -14.1171
8.0149 | -34.4428 | -14.1272
8.021 | -34.4432 | -14.1374
8.0271 | -34.4436 | -14.1478
8.0332 | -34.4439 | -14.1579
8.0393 | -34.4444 | -14.1682
8.0454 | -34.4447 | -14.1786
8.0515 | -34.4452 | -14.1887
8.0576 | -34.4455 | -14.1991
8.0637 | -34.4459 | -14.2091
8.0698 | -34.4463 | -14.2191
8.0759 | -34.4466 | -14.2292
8.082 | -34.447 | -14.2392
8.0881 | -34.4474 | -14.2494
8.0942 | -34.4478 | -14.2596
8.1004 | -34.4482 | -14.2699
8.1065 | -34.4486 | -14.2799
8.1126 | -34.4489 | -14.29
8.1187 | -34.4493 | -14.3003
8.1248 | -34.4497 | -14.3102
8.1309 | -34.4501 | -14.3202
8.137 | -34.4505 | -14.3302
8.1431 | -34.4509 | -14.3405
8.1492 | -34.4512 | -14.3506
8.1553 | -34.4516 | -14.3608
8.1614 | -34.452 | -14.3711
8.1675 | -34.4524 | -14.3812
8.1736 | -34.4527 | -14.3916
8.1797 | -34.4531 | -14.4018
8.1858 | -34.4535 | -14.412
8.1919 | -34.4539 | -14.4221
8.198 | -34.4543 | -14.4321
8.2041 | -34.4547 | -14.4423
8.2102 | -34.4551 | -14.4524
8.2163 | -34.4555 | -14.4624
8.2224 | -34.4558 | -14.4725
8.2285 | -34.4562 | -14.4824
8.2346 | -34.4566 | -14.4925
8.2408 | -34.457 | -14.5024
8.2469 | -34.4574 | -14.5126
8.253 | -34.4578 | -14.5226
8.2591 | -34.4582 | -14.5328
8.2652 | -34.4586 | -14.543
8.2713 | -34.459 | -14.5532
8.2774 | -34.4594 | -14.5634
8.2835 | -34.4598 | -14.5738
8.2896 | -34.4602 | -14.5838
8.2957 | -34.4606 | -14.594
8.3018 | -34.4609 | -14.6041
8.3079 | -34.4613 | -14.6141
8.314 | -34.4617 | -14.6242
8.3201 | -34.4622 | -14.6342
8.3262 | -34.4625 | -14.6444
8.3323 | -34.4629 | -14.6545
8.3384 | -34.4633 | -14.6648
8.3445 | -34.4637 | -14.6748
8.3506 | -34.4641 | -14.685
8.3567 | -34.4645 | -14.6952
8.3628 | -34.4649 | -14.7053
8.3689 | -34.4652 | -14.7153
8.375 | -34.4656 | -14.7255
8.3812 | -34.466 | -14.7355
8.3873 | -34.4664 | -14.7457
8.3934 | -34.4668 | -14.7559
8.3995 | -34.4672 | -14.7659
8.4056 | -34.4676 | -14.776
8.4117 | -34.4679 | -14.7862
8.4178 | -34.4683 | -14.7963
8.4239 | -34.4687 | -14.8063
8.43 | -34.4691 | -14.8164
8.4361 | -34.4694 | -14.8266
8.4422 | -34.4698 | -14.8366
8.4483 | -34.4702 | -14.8466
8.4544 | -34.4706 | -14.8567
8.4605 | -34.471 | -14.8667
8.4666 | -34.4714 | -14.877
8.4727 | -34.4718 | -14.8872
8.4788 | -34.4721 | -14.8972
8.4849 | -34.4725 | -14.9073
8.491 | -34.4729 | -14.9174
8.4971 | -34.4733 | -14.9273
8.5032 | -34.4737 | -14.9372
8.5093 | -34.4741 | -14.9472
8.5154 | -34.4745 | -14.9573
8.5215 | -34.4749 | -14.9673
8.5277 | -34.4752 | -14.9774
8.5338 | -34.4756 | -14.9873
8.5399 | -34.476 | -14.9974
8.546 | -34.4764 | -15.0073
8.5521 | -34.4768 | -15.0174
8.5582 | -34.4772 | -15.0274
8.5643 | -34.4776 | -15.0372
8.5704 | -34.478 | -15.0469
8.5765 | -34.4784 | -15.0571
8.5826 | -34.4788 | -15.0668
8.5887 | -34.4792 | -15.0768
8.5948 | -34.4796 | -15.0869
8.6009 | -34.48 | -15.0967
8.607 | -34.4803 | -15.1069
8.6131 | -34.4807 | -15.117
8.6192 | -34.4812 | -15.1268
8.6253 | -34.4816 | -15.1368
8.6314 | -34.4819 | -15.1467
8.6375 | -34.4823 | -15.1568
8.6436 | -34.4827 | -15.1668
8.6497 | -34.4831 | -15.1766
8.6558 | -34.4834 | -15.1866
8.6619 | -34.4838 | -15.1966
8.6681 | -34.4842 | -15.2068
8.6742 | -34.4846 | -15.2167
8.6803 | -34.485 | -15.2267
8.6864 | -34.4854 | -15.2364
8.6925 | -34.4858 | -15.2462
8.6986 | -34.4861 | -15.2563
8.7047 | -34.4865 | -15.2664
8.7108 | -34.4869 | -15.2764
8.7169 | -34.4873 | -15.2864
8.723 | -34.4877 | -15.2961
8.7291 | -34.4881 | -15.3059
8.7352 | -34.4885 | -15.3161
8.7413 | -34.4888 | -15.3261
8.7474 | -34.4892 | -15.3359
8.7535 | -34.4896 | -15.3459
8.7596 | -34.4901 | -15.3559
8.7657 | -34.4904 | -15.3659
8.7718 | -34.4908 | -15.3759
8.7779 | -34.4912 | -15.386
8.784 | -34.4916 | -15.396
8.7901 | -34.492 | -15.4059
8.7962 | -34.4924 | -15.4156
8.8023 | -34.4928 | -15.4255
8.8084 | -34.4932 | -15.4354
8.8146 | -34.4936 | -15.4453
8.8207 | -34.4939 | -15.4551
8.8268 | -34.4944 | -15.4647
8.8329 | -34.4947 | -15.4746
8.839 | -34.4952 | -15.4846
8.8451 | -34.4955 | -15.4944
8.8512 | -34.496 | -15.5043
8.8573 | -34.4963 | -15.5142
8.8634 | -34.4967 | -15.5239
8.8695 | -34.4971 | -15.534
8.8756 | -34.4975 | -15.5439
8.8817 | -34.4979 | -15.5536
8.8878 | -34.4983 | -15.5638
8.8939 | -34.4987 | -15.5736
8.9 | -34.4991 | -15.5835
8.9061 | -34.4994 | -15.5934
8.9122 | -34.4999 | -15.6033
8.9183 | -34.5003 | -15.6132
8.9244 | -34.5007 | -15.6232
8.9305 | -34.5011 | -15.6328
8.9366 | -34.5015 | -15.6427
8.9427 | -34.5019 | -15.6527
8.9488 | -34.5023 | -15.6625
8.955 | -34.5027 | -15.6724
8.9611 | -34.5031 | -15.6824
8.9672 | -34.5035 | -15.6922
8.9733 | -34.5039 | -15.7019
8.9794 | -34.5044 | -15.7116
8.9855 | -34.5048 | -15.7216
8.9916 | -34.5051 | -15.7316
8.9977 | -34.5055 | -15.7415
9.0038 | -34.506 | -15.7512
9.0099 | -34.5064 | -15.7611
9.016 | -34.5068 | -15.771
9.0221 | -34.5072 | -15.7807
9.0282 | -34.5076 | -15.7904
9.0343 | -34.508 | -15.8003
9.0404 | -34.5085 | -15.8101
9.0465 | -34.5089 | -15.8202
9.0526 | -34.5093 | -15.8299
9.0587 | -34.5097 | -15.8398
9.0648 | -34.5101 | -15.8495
9.0709 | -34.5105 | -15.8593
9.077 | -34.511 | -15.8693
9.0831 | -34.5114 | -15.8793
9.0892 | -34.5118 | -15.8892
9.0953 | -34.5122 | -15.8991
9.1015 | -34.5126 | -15.9089
9.1076 | -34.5131 | -15.9185
9.1137 | -34.5135 | -15.9283
9.1198 | -34.5139 | -15.9381
9.1259 | -34.5143 | -15.948
9.132 | -34.5147 | -15.958
9.1381 | -34.5151 | -15.9677
9.1442 | -34.5156 | -15.9771
9.1503 | -34.516 | -15.9871
9.1564 | -34.5164 | -15.9969
9.1625 | -34.5168 | -16.0067
9.1686 | -34.5173 | -16.0164
9.1747 | -34.5177 | -16.0262
9.1808 | -34.5181 | -16.0359
9.1869 | -34.5186 | -16.0457
9.193 | -34.519 | -16.0556
9.1991 | -34.5194 | -16.0654
9.2052 | -34.5198 | -16.0753
9.2113 | -34.5202 | -16.085
9.2174 | -34.5207 | -16.0948
9.2235 | -34.5211 | -16.1046
9.2296 | -34.5216 | -16.1143
9.2357 | -34.522 | -16.1243
9.2419 | -34.5224 | -16.134
9.248 | -34.5229 | -16.1437
9.2541 | -34.5233 | -16.1536
9.2602 | -34.5238 | -16.1635
9.2663 | -34.5242 | -16.1731
9.2724 | -34.5246 | -16.1828
9.2785 | -34.525 | -16.1925
9.2846 | -34.5255 | -16.2022
9.2907 | -34.5259 | -16.212
9.2968 | -34.5264 | -16.2219
9.3029 | -34.5268 | -16.2318
9.309 | -34.5272 | -16.2416
9.3151 | -34.5276 | -16.2513
9.3212 | -34.5281 | -16.2609
9.3273 | -34.5285 | -16.2706
9.3334 | -34.5289 | -16.2803
9.3395 | -34.5294 | -16.2899
9.3456 | -34.5298 | -16.2995
9.3517 | -34.5302 | -16.3094
9.3578 | -34.5306 | -16.3191
9.3639 | -34.5311 | -16.3287
9.37 | -34.5315 | -16.3385
9.3761 | -34.532 | -16.3484
9.3822 | -34.5324 | -16.3582
9.3884 | -34.5329 | -16.3677
9.3945 | -34.5333 | -16.3774
9.4006 | -34.5338 | -16.3872
9.4067 | -34.5342 | -16.3971
9.4128 | -34.5346 | -16.4069
9.4189 | -34.5351 | -16.4169
9.425 | -34.5355 | -16.4265
9.4311 | -34.536 | -16.4363
9.4372 | -34.5364 | -16.446
9.4433 | -34.5369 | -16.4557
9.4494 | -34.5373 | -16.4653
9.4555 | -34.5378 | -16.4751
9.4616 | -34.5382 | -16.4848
9.4677 | -34.5386 | -16.4945
9.4738 | -34.5391 | -16.5043
9.4799 | -34.5395 | -16.514
9.486 | -34.5399 | -16.5238
9.4921 | -34.5404 | -16.5335
9.4982 | -34.5409 | -16.5431
9.5043 | -34.5413 | -16.5527
9.5104 | -34.5417 | -16.5623
9.5165 | -34.5422 | -16.572
9.5226 | -34.5426 | -16.5818
9.5288 | -34.5431 | -16.5913
9.5349 | -34.5435 | -16.6011
9.541 | -34.5439 | -16.6106
9.5471 | -34.5444 | -16.6203
9.5532 | -34.5448 | -16.6301
9.5593 | -34.5453 | -16.6399
9.5654 | -34.5457 | -16.6495
9.5715 | -34.5462 | -16.6592
9.5776 | -34.5466 | -16.6691
9.5837 | -34.547 | -16.6788
9.5898 | -34.5474 | -16.6886
9.5959 | -34.5479 | -16.6983
9.602 | -34.5483 | -16.7079
9.6081 | -34.5488 | -16.7175
9.6142 | -34.5492 | -16.7272
9.6203 | -34.5497 | -16.7369
9.6264 | -34.5501 | -16.7464
9.6325 | -34.5505 | -16.7562
9.6386 | -34.551 | -16.7657
9.6447 | -34.5514 | -16.7754
9.6508 | -34.5518 | -16.785
9.6569 | -34.5523 | -16.7947
9.663 | -34.5527 | -16.8045
9.6691 | -34.5532 | -16.8141
9.6753 | -34.5536 | -16.8238
9.6814 | -34.5541 | -16.8334
9.6875 | -34.5545 | -16.8429
9.6936 | -34.555 | -16.8524
9.6997 | -34.5554 | -16.8622
9.7058 | -34.5559 | -16.8718
9.7119 | -34.5564 | -16.8815
9.718 | -34.5568 | -16.8913
9.7241 | -34.5572 | -16.9009
9.7302 | -34.5577 | -16.9106
9.7363 | -34.5581 | -16.9201
9.7424 | -34.5586 | -16.9297
9.7485 | -34.559 | -16.9394
9.7546 | -34.5594 | -16.949
9.7607 | -34.5599 | -16.9588
9.7668 | -34.5603 | -16.9685
9.7729 | -34.5608 | -16.9782
9.779 | -34.5613 | -16.9876
9.7851 | -34.5617 | -16.9974
9.7912 | -34.5622 | -17.007
9.7973 | -34.5627 | -17.0168
9.8034 | -34.5632 | -17.0263
9.8095 | -34.5636 | -17.0362
9.8157 | -34.5641 | -17.0459
9.8218 | -34.5645 | -17.0558
9.8279 | -34.565 | -17.0653
9.834 | -34.5654 | -17.075
9.8401 | -34.5659 | -17.0847
9.8462 | -34.5663 | -17.0941
9.8523 | -34.5668 | -17.1037
9.8584 | -34.5672 | -17.1133
9.8645 | -34.5676 | -17.123
9.8706 | -34.5681 | -17.1327
9.8767 | -34.5685 | -17.1423
9.8828 | -34.5689 | -17.1519
9.8889 | -34.5694 | -17.1614
9.895 | -34.5698 | -17.171
9.9011 | -34.5703 | -17.1807
9.9072 | -34.5708 | -17.1902
9.9133 | -34.5712 | -17.1998
9.9194 | -34.5716 | -17.2092
9.9255 | -34.5721 | -17.2189
9.9316 | -34.5725 | -17.2284
9.9377 | -34.573 | -17.238
9.9438 | -34.5734 | -17.2476
9.9499 | -34.5739 | -17.2575
9.956 | -34.5743 | -17.2671
9.9622 | -34.5748 | -17.2769
9.9683 | -34.5752 | -17.2865
9.9744 | -34.5757 | -17.2962
9.9805 | -34.5761 | -17.3058
9.9866 | -34.5766 | -17.3154
9.9927 | -34.577 | -17.3249
9.9988 | -34.5775 | -17.3346
10.0049 | -34.5779 | -17.3442
10.011 | -34.5784 | -17.3539
10.0171 | -34.5788 | -17.3634
10.0232 | -34.5793 | -17.3729
10.0293 | -34.5798 | -17.3825
10.0354 | -34.5802 | -17.3921
10.0415 | -34.5807 | -17.4017
10.0476 | -34.5811 | -17.4114
10.0537 | -34.5816 | -17.4212
10.0598 | -34.582 | -17.4308
10.0659 | -34.5824 | -17.4404
10.072 | -34.5829 | -17.4498
10.0781 | -34.5833 | -17.4594
10.0842 | -34.5838 | -17.4689
10.0903 | -34.5843 | -17.4786
10.0964 | -34.5847 | -17.488
10.1026 | -34.5852 | -17.4976
10.1087 | -34.5856 | -17.5072
10.1148 | -34.586 | -17.5169
10.1209 | -34.5865 | -17.5263
10.127 | -34.5869 | -17.5357
10.1331 | -34.5874 | -17.5452
10.1392 | -34.5878 | -17.5547
10.1453 | -34.5883 | -17.5643
10.1514 | -34.5887 | -17.574
10.1575 | -34.5892 | -17.5836
10.1636 | -34.5897 | -17.5928
10.1697 | -34.5901 | -17.6024
10.1758 | -34.5906 | -17.6119
10.1819 | -34.591 | -17.6214
10.188 | -34.5915 | -17.631
10.1941 | -34.5919 | -17.6405
10.2002 | -34.5923 | -17.65
10.2063 | -34.5928 | -17.6595
10.2124 | -34.5933 | -17.6691
10.2185 | -34.5937 | -17.6787
10.2246 | -34.5942 | -17.6881
10.2307 | -34.5946 | -17.6978
10.2368 | -34.595 | -17.707
10.2429 | -34.5955 | -17.7165
10.2491 | -34.5959 | -17.7262
10.2552 | -34.5964 | -17.7357
10.2613 | -34.5968 | -17.7453
10.2674 | -34.5973 | -17.7547
10.2735 | -34.5977 | -17.7645
10.2796 | -34.5982 | -17.7739
10.2857 | -34.5987 | -17.7835
10.2918 | -34.5991 | -17.7933
10.2979 | -34.5996 | -17.8027
10.304 | -34.6 | -17.8121
10.3101 | -34.6004 | -17.8218
10.3162 | -34.6009 | -17.8313
10.3223 | -34.6014 | -17.8407
10.3284 | -34.6018 | -17.8503
10.3345 | -34.6022 | -17.8599
10.3406 | -34.6027 | -17.8694
10.3467 | -34.6031 | -17.8787
10.3528 | -34.6035 | -17.8881
10.3589 | -34.604 | -17.8975
10.365 | -34.6044 | -17.9074
10.3711 | -34.6049 | -17.917
10.3772 | -34.6053 | -17.9263
10.3833 | -34.6058 | -17.9359
10.3895 | -34.6062 | -17.9454
10.3956 | -34.6067 | -17.9549
10.4017 | -34.6072 | -17.9642
10.4078 | -34.6076 | -17.9737
10.4139 | -34.608 | -17.9831
10.42 | -34.6085 | -17.9929
10.4261 | -34.6089 | -18.0021
10.4322 | -34.6094 | -18.0117
10.4383 | -34.6098 | -18.0211
10.4444 | -34.6103 | -18.0307
10.4505 | -34.6107 | -18.0401
10.4566 | -34.6111 | -18.0493
10.4627 | -34.6116 | -18.0589
10.4688 | -34.612 | -18.0683
10.4749 | -34.6125 | -18.0779
10.481 | -34.6129 | -18.0874
10.4871 | -34.6134 | -18.0968
10.4932 | -34.6139 | -18.1061
10.4993 | -34.6143 | -18.1154
10.5054 | -34.6148 | -18.1248
10.5115 | -34.6153 | -18.1344
10.5176 | -34.6157 | -18.1437
10.5237 | -34.6162 | -18.1534
10.5298 | -34.6166 | -18.1627
10.536 | -34.6171 | -18.1723
10.5421 | -34.6175 | -18.1817
10.5482 | -34.6179 | -18.1912
10.5543 | -34.6184 | -18.2005
10.5604 | -34.6189 | -18.2098
10.5665 | -34.6194 | -18.219
10.5726 | -34.6198 | -18.2283
10.5787 | -34.6202 | -18.2374
10.5848 | -34.6206 | -18.2465
10.5909 | -34.6211 | -18.256
10.597 | -34.6215 | -18.2653
10.6031 | -34.622 | -18.2751
10.6092 | -34.6224 | -18.2846
10.6153 | -34.6229 | -18.294
10.6214 | -34.6234 | -18.3032
10.6275 | -34.6238 | -18.3125
10.6336 | -34.6243 | -18.322
10.6397 | -34.6247 | -18.3314
10.6458 | -34.6252 | -18.3409
10.6519 | -34.6257 | -18.3505
10.658 | -34.6261 | -18.3602
10.6641 | -34.6265 | -18.3694
10.6702 | -34.627 | -18.3789
10.6764 | -34.6274 | -18.3884
10.6825 | -34.6279 | -18.3977
10.6886 | -34.6283 | -18.4071
10.6947 | -34.6288 | -18.4164
10.7008 | -34.6292 | -18.4261
10.7069 | -34.6297 | -18.4356
10.713 | -34.6301 | -18.4448
10.7191 | -34.6305 | -18.454
10.7252 | -34.6309 | -18.4629
10.7313 | -34.6314 | -18.4724
10.7374 | -34.6318 | -18.4819
10.7435 | -34.6322 | -18.4915
10.7496 | -34.6327 | -18.5008
10.7557 | -34.6331 | -18.5102
10.7618 | -34.6336 | -18.5193
10.7679 | -34.634 | -18.5286
10.774 | -34.6345 | -18.538
10.7801 | -34.6349 | -18.5476
10.7862 | -34.6354 | -18.5568
10.7923 | -34.6358 | -18.5664
10.7984 | -34.6362 | -18.5758
10.8045 | -34.6367 | -18.585
10.8106 | -34.6372 | -18.5947
10.8168 | -34.6376 | -18.6043
10.8229 | -34.638 | -18.6137
10.829 | -34.6384 | -18.6228
10.8351 | -34.639 | -18.6325
10.8412 | -34.6395 | -18.6414
10.8473 | -34.6399 | -18.651
10.8534 | -34.6404 | -18.6606
10.8595 | -34.6408 | -18.6701
10.8656 | -34.6412 | -18.6795
10.8717 | -34.6416 | -18.6892
10.8778 | -34.6421 | -18.6981
10.8839 | -34.6426 | -18.7073
10.89 | -34.643 | -18.7166
10.8961 | -34.6435 | -18.7262
10.9022 | -34.6439 | -18.7355
10.9083 | -34.6444 | -18.7449
10.9144 | -34.6448 | -18.7542
10.9205 | -34.6453 | -18.7637
10.9266 | -34.6458 | -18.7728
10.9327 | -34.6462 | -18.7823
10.9388 | -34.6467 | -18.7915
10.9449 | -34.6471 | -18.8009
10.951 | -34.6475 | -18.8103
10.9571 | -34.6479 | -18.8198
10.9633 | -34.6484 | -18.8291
10.9694 | -34.6489 | -18.8385
10.9755 | -34.6493 | -18.8478
10.9816 | -34.6498 | -18.8571
10.9877 | -34.6503 | -18.8662
10.9938 | -34.6508 | -18.8757
10.9999 | -34.6512 | -18.8852
11.006 | -34.6516 | -18.8945
11.0121 | -34.6521 | -18.9041
11.0182 | -34.6526 | -18.9135
11.0243 | -34.6531 | -18.9229
11.0304 | -34.6535 | -18.9321
11.0365 | -34.6539 | -18.9412
11.0426 | -34.6545 | -18.9507
11.0487 | -34.655 | -18.96
11.0548 | -34.6555 | -18.9693
11.0609 | -34.6559 | -18.9788
11.067 | -34.6564 | -18.9883
11.0731 | -34.6569 | -18.9975
11.0792 | -34.6573 | -19.0071
11.0853 | -34.6578 | -19.0166
11.0914 | -34.6583 | -19.0262
11.0975 | -34.6587 | -19.0354
11.1037 | -34.6591 | -19.0444
11.1098 | -34.6596 | -19.0541
11.1159 | -34.66 | -19.0633
11.122 | -34.6605 | -19.0721
11.1281 | -34.6609 | -19.0813
11.1342 | -34.6614 | -19.0904
11.1403 | -34.6619 | -19.0998
11.1464 | -34.6624 | -19.1091
11.1525 | -34.6628 | -19.1183
11.1586 | -34.6633 | -19.1276
11.1647 | -34.6637 | -19.1366
11.1708 | -34.6642 | -19.146
11.1769 | -34.6646 | -19.1555
11.183 | -34.6652 | -19.1646
11.1891 | -34.6656 | -19.174
11.1952 | -34.6662 | -19.1832
11.2013 | -34.6666 | -19.1929
11.2074 | -34.6671 | -19.2023
11.2135 | -34.6676 | -19.2121
11.2196 | -34.6681 | -19.2218
11.2257 | -34.6686 | -19.2315
11.2318 | -34.6691 | -19.2406
11.2379 | -34.6695 | -19.2499
11.244 | -34.67 | -19.2594
11.2502 | -34.6705 | -19.2688
11.2563 | -34.671 | -19.2781
11.2624 | -34.6714 | -19.2874
11.2685 | -34.6719 | -19.2968
11.2746 | -34.6723 | -19.3061
11.2807 | -34.6728 | -19.3157
11.2868 | -34.6733 | -19.325
11.2929 | -34.6737 | -19.334
11.299 | -34.6742 | -19.3434
11.3051 | -34.6747 | -19.353
11.3112 | -34.6752 | -19.3623
11.3173 | -34.6756 | -19.3713
11.3234 | -34.6762 | -19.3803
11.3295 | -34.6767 | -19.3895
11.3356 | -34.6771 | -19.3989
11.3417 | -34.6776 | -19.4079
11.3478 | -34.6781 | -19.4173
11.3539 | -34.6785 | -19.4265
11.36 | -34.6789 | -19.4357
11.3661 | -34.6794 | -19.4449
11.3722 | -34.6799 | -19.454
11.3783 | -34.6804 | -19.4633
11.3844 | -34.6809 | -19.4726
11.3906 | -34.6814 | -19.482
11.3967 | -34.6819 | -19.4907
11.4028 | -34.6824 | -19.5004
11.4089 | -34.6828 | -19.5097
11.415 | -34.6833 | -19.519
11.4211 | -34.6837 | -19.5286
11.4272 | -34.6842 | -19.5379
11.4333 | -34.6846 | -19.5469
11.4394 | -34.6851 | -19.556
11.4455 | -34.6855 | -19.5651
11.4516 | -34.686 | -19.5743
11.4577 | -34.6865 | -19.5834
11.4638 | -34.687 | -19.5923
11.4699 | -34.6875 | -19.6015
11.476 | -34.6879 | -19.6112
11.4821 | -34.6884 | -19.6206
11.4882 | -34.6889 | -19.6301
11.4943 | -34.6893 | -19.6396
11.5004 | -34.6898 | -19.6485
11.5065 | -34.6903 | -19.6577
11.5126 | -34.6908 | -19.6668
11.5187 | -34.6912 | -19.6762
11.5248 | -34.6917 | -19.6857
11.5309 | -34.6922 | -19.695
11.5371 | -34.6927 | -19.7043
11.5432 | -34.6932 | -19.7135
11.5493 | -34.6936 | -19.7231
11.5554 | -34.6941 | -19.7325
11.5615 | -34.6946 | -19.7419
11.5676 | -34.6951 | -19.7513
11.5737 | -34.6956 | -19.7608
11.5798 | -34.6961 | -19.77
11.5859 | -34.6966 | -19.7793
11.592 | -34.6971 | -19.7887
11.5981 | -34.6976 | -19.7979
11.6042 | -34.6981 | -19.8074
11.6103 | -34.6985 | -19.8166
11.6164 | -34.699 | -19.8259
11.6225 | -34.6995 | -19.8349
11.6286 | -34.7 | -19.8441
11.6347 | -34.7004 | -19.8533
11.6408 | -34.701 | -19.8627
11.6469 | -34.7014 | -19.8718
11.653 | -34.7019 | -19.8809
11.6591 | -34.7024 | -19.89
11.6652 | -34.7028 | -19.8991
11.6713 | -34.7033 | -19.9083
11.6775 | -34.7038 | -19.9176
11.6836 | -34.7043 | -19.9269
11.6897 | -34.7048 | -19.9363
11.6958 | -34.7053 | -19.9457
11.7019 | -34.7058 | -19.9551
11.708 | -34.7063 | -19.9643
11.7141 | -34.7068 | -19.9735
11.7202 | -34.7073 | -19.9828
11.7263 | -34.7077 | -19.9922
11.7324 | -34.7082 | -20.0015
11.7385 | -34.7087 | -20.0109
11.7446 | -34.7091 | -20.0202
11.7507 | -34.7097 | -20.0293
11.7568 | -34.7101 | -20.0385
11.7629 | -34.7106 | -20.0478
11.769 | -34.7111 | -20.0569
11.7751 | -34.7115 | -20.0662
11.7812 | -34.7121 | -20.0752
11.7873 | -34.7126 | -20.0847
11.7934 | -34.713 | -20.0939
11.7995 | -34.7135 | -20.1031
11.8056 | -34.714 | -20.1123
11.8117 | -34.7145 | -20.1215
11.8178 | -34.715 | -20.1306
11.824 | -34.7155 | -20.1398
11.8301 | -34.716 | -20.1489
11.8362 | -34.7165 | -20.1581
11.8423 | -34.717 | -20.1673
11.8484 | -34.7174 | -20.1767
11.8545 | -34.718 | -20.1857
11.8606 | -34.7185 | -20.1949
11.8667 | -34.7189 | -20.2043
11.8728 | -34.7194 | -20.2132
11.8789 | -34.7199 | -20.2225
11.885 | -34.7204 | -20.2318
11.8911 | -34.721 | -20.241
11.8972 | -34.7215 | -20.2501
11.9033 | -34.7219 | -20.2591
11.9094 | -34.7224 | -20.2684
11.9155 | -34.7229 | -20.2776
11.9216 | -34.7234 | -20.2866
11.9277 | -34.7239 | -20.2959
11.9338 | -34.7244 | -20.305
11.9399 | -34.7249 | -20.3144
11.946 | -34.7254 | -20.3235
11.9521 | -34.7259 | -20.3326
11.9582 | -34.7264 | -20.3418
11.9644 | -34.7268 | -20.3507
11.9705 | -34.7274 | -20.36
11.9766 | -34.7279 | -20.3691
11.9827 | -34.7284 | -20.3783
11.9888 | -34.7289 | -20.3873
11.9949 | -34.7294 | -20.3964
12.001 | -34.7299 | -20.4057
12.0071 | -34.7304 | -20.4147
12.0132 | -34.7309 | -20.4238
12.0193 | -34.7313 | -20.433
12.0254 | -34.7318 | -20.4421
12.0315 | -34.7323 | -20.4511
12.0376 | -34.7328 | -20.4603
12.0437 | -34.7333 | -20.4694
12.0498 | -34.7338 | -20.4786
12.0559 | -34.7343 | -20.4877
12.062 | -34.7348 | -20.4968
12.0681 | -34.7353 | -20.5059
12.0742 | -34.7358 | -20.515
12.0803 | -34.7363 | -20.5243
12.0864 | -34.7368 | -20.5334
12.0925 | -34.7373 | -20.5425
12.0986 | -34.7378 | -20.5517
12.1047 | -34.7383 | -20.5609
12.1109 | -34.7388 | -20.5699
12.117 | -34.7393 | -20.579
12.1231 | -34.7398 | -20.5882
12.1292 | -34.7402 | -20.5973
12.1353 | -34.7408 | -20.6062
12.1414 | -34.7413 | -20.6154
12.1475 | -34.7418 | -20.6245
12.1536 | -34.7423 | -20.6336
12.1597 | -34.7428 | -20.6426
12.1658 | -34.7433 | -20.6515
12.1719 | -34.7438 | -20.6604
12.178 | -34.7442 | -20.6696
12.1841 | -34.7448 | -20.6788
12.1902 | -34.7453 | -20.6879
12.1963 | -34.7458 | -20.6971
12.2024 | -34.7463 | -20.7061
12.2085 | -34.7468 | -20.7152
12.2146 | -34.7473 | -20.7245
12.2207 | -34.7478 | -20.7336
12.2268 | -34.7483 | -20.7428
12.2329 | -34.7488 | -20.7517
12.239 | -34.7493 | -20.7605
12.2451 | -34.7498 | -20.7696
12.2513 | -34.7503 | -20.7788
12.2574 | -34.7508 | -20.7877
12.2635 | -34.7513 | -20.7968
12.2696 | -34.7518 | -20.8058
12.2757 | -34.7523 | -20.8147
12.2818 | -34.7528 | -20.8239
12.2879 | -34.7533 | -20.8329
12.294 | -34.7538 | -20.8419
12.3001 | -34.7543 | -20.851
12.3062 | -34.7548 | -20.8601
12.3123 | -34.7554 | -20.8691
12.3184 | -34.7559 | -20.8781
12.3245 | -34.7564 | -20.887
12.3306 | -34.7569 | -20.8962
12.3367 | -34.7574 | -20.9052
12.3428 | -34.7579 | -20.9143
12.3489 | -34.7584 | -20.9231
12.355 | -34.7589 | -20.9323
12.3611 | -34.7594 | -20.9415
12.3672 | -34.7599 | -20.9504
12.3733 | -34.7604 | -20.9595
12.3794 | -34.7609 | -20.9683
12.3855 | -34.7614 | -20.9773
12.3916 | -34.7619 | -20.9863
12.3978 | -34.7624 | -20.9954
12.4039 | -34.763 | -21.0045
12.41 | -34.7634 | -21.0137
12.4161 | -34.764 | -21.0222
12.4222 | -34.7644 | -21.031
12.4283 | -34.7649 | -21.04
12.4344 | -34.7655 | -21.0491
12.4405 | -34.7659 | -21.0582
12.4466 | -34.7664 | -21.0672
12.4527 | -34.7669 | -21.076
12.4588 | -34.7674 | -21.085
12.4649 | -34.7679 | -21.0939
12.471 | -34.7685 | -21.1032
12.4771 | -34.7689 | -21.1122
12.4832 | -34.7695 | -21.1211
12.4893 | -34.7699 | -21.1299
12.4954 | -34.7704 | -21.1389
12.5015 | -34.7709 | -21.1476
12.5076 | -34.7714 | -21.1566
12.5137 | -34.7719 | -21.1656
12.5198 | -34.7724 | -21.1745
12.5259 | -34.7729 | -21.1834
12.532 | -34.7734 | -21.1925
12.5382 | -34.7739 | -21.2014
12.5443 | -34.7744 | -21.2102
12.5504 | -34.7748 | -21.2193
12.5565 | -34.7754 | -21.2282
12.5626 | -34.7759 | -21.2375
12.5687 | -34.7764 | -21.2463
12.5748 | -34.7769 | -21.2552
12.5809 | -34.7773 | -21.2642
12.587 | -34.7778 | -21.2732
12.5931 | -34.7784 | -21.2818
12.5992 | -34.7789 | -21.2908
12.6053 | -34.7794 | -21.2998
12.6114 | -34.7799 | -21.3087
12.6175 | -34.7804 | -21.3177
12.6236 | -34.7809 | -21.3268
12.6297 | -34.7814 | -21.3359
12.6358 | -34.7819 | -21.3446
12.6419 | -34.7824 | -21.3535
12.648 | -34.7829 | -21.3627
12.6541 | -34.7834 | -21.3716
12.6602 | -34.7839 | -21.38
12.6663 | -34.7844 | -21.3889
12.6724 | -34.7849 | -21.3975
12.6785 | -34.7853 | -21.4063
12.6847 | -34.7858 | -21.4149
12.6908 | -34.7863 | -21.4238
12.6969 | -34.7869 | -21.4327
12.703 | -34.7873 | -21.4416
12.7091 | -34.7878 | -21.4505
12.7152 | -34.7883 | -21.4592
12.7213 | -34.7888 | -21.468
12.7274 | -34.7893 | -21.4769
12.7335 | -34.7898 | -21.4856
12.7396 | -34.7903 | -21.4944
12.7457 | -34.7909 | -21.5034
12.7518 | -34.7913 | -21.512
12.7579 | -34.7918 | -21.5209
12.764 | -34.7923 | -21.5299
12.7701 | -34.7928 | -21.5385
12.7762 | -34.7933 | -21.5473
12.7823 | -34.7938 | -21.5562
12.7884 | -34.7943 | -21.5653
12.7945 | -34.7948 | -21.5743
12.8006 | -34.7953 | -21.5832
12.8067 | -34.7958 | -21.5921
12.8128 | -34.7963 | -21.6007
12.8189 | -34.7968 | -21.6096
12.8251 | -34.7973 | -21.6186
12.8312 | -34.7978 | -21.6276
12.8373 | -34.7983 | -21.6362
12.8434 | -34.7988 | -21.6451
12.8495 | -34.7993 | -21.6537
12.8556 | -34.7997 | -21.6626
12.8617 | -34.8002 | -21.6715
12.8678 | -34.8007 | -21.6804
12.8739 | -34.8012 | -21.6888
12.88 | -34.8017 | -21.6976
12.8861 | -34.8022 | -21.7064
12.8922 | -34.8027 | -21.7154
12.8983 | -34.8032 | -21.7242
12.9044 | -34.8037 | -21.7329
12.9105 | -34.8041 | -21.7414
12.9166 | -34.8046 | -21.7502
12.9227 | -34.8051 | -21.759
12.9288 | -34.8056 | -21.7677
12.9349 | -34.8061 | -21.7765
12.941 | -34.8066 | -21.7851
12.9471 | -34.8071 | -21.794
12.9532 | -34.8077 | -21.8028
12.9593 | -34.8082 | -21.8117
12.9654 | -34.8087 | -21.8206
12.9716 | -34.8092 | -21.8296
12.9777 | -34.8097 | -21.8382
12.9838 | -34.8102 | -21.8469
12.9899 | -34.8107 | -21.8558
12.996 | -34.8112 | -21.8643
13.0021 | -34.8117 | -21.873
13.0082 | -34.8122 | -21.8816
13.0143 | -34.8127 | -21.8901
13.0204 | -34.8132 | -21.8988
13.0265 | -34.8137 | -21.9077
13.0326 | -34.8142 | -21.9165
13.0387 | -34.8147 | -21.9252
13.0448 | -34.8152 | -21.934
13.0509 | -34.8157 | -21.9424
13.057 | -34.8162 | -21.951
13.0631 | -34.8167 | -21.9596
13.0692 | -34.8172 | -21.9683
13.0753 | -34.8177 | -21.9772
13.0814 | -34.8181 | -21.986
13.0875 | -34.8186 | -21.9943
13.0936 | -34.8191 | -22.0031
13.0997 | -34.8196 | -22.0118
13.1058 | -34.8201 | -22.0203
13.112 | -34.8206 | -22.0291
13.1181 | -34.8211 | -22.0373
13.1242 | -34.8216 | -22.0459
13.1303 | -34.8221 | -22.0547
13.1364 | -34.8226 | -22.0636
13.1425 | -34.823 | -22.0722
13.1486 | -34.8235 | -22.081
13.1547 | -34.824 | -22.0894
13.1608 | -34.8245 | -22.0982
13.1669 | -34.8249 | -22.1069
13.173 | -34.8255 | -22.1157
13.1791 | -34.8259 | -22.1243
13.1852 | -34.8264 | -22.1331
13.1913 | -34.8269 | -22.1416
13.1974 | -34.8274 | -22.1503
13.2035 | -34.8278 | -22.1589
13.2096 | -34.8283 | -22.1673
13.2157 | -34.8288 | -22.176
13.2218 | -34.8293 | -22.1846
13.2279 | -34.8298 | -22.1927
13.234 | -34.8303 | -22.2013
13.2401 | -34.8308 | -22.2098
13.2462 | -34.8312 | -22.2184
13.2524 | -34.8317 | -22.227
13.2585 | -34.8323 | -22.2356
13.2646 | -34.8327 | -22.2443
13.2707 | -34.8332 | -22.2529
13.2768 | -34.8337 | -22.2619
13.2829 | -34.8342 | -22.2707
13.289 | -34.8347 | -22.2793
13.2951 | -34.8352 | -22.2876
13.3012 | -34.8357 | -22.2962
13.3073 | -34.8362 | -22.3051
13.3134 | -34.8366 | -22.3138
13.3195 | -34.8371 | -22.3226
13.3256 | -34.8376 | -22.3311
13.3317 | -34.8381 | -22.3396
13.3378 | -34.8386 | -22.3484
13.3439 | -34.8391 | -22.3572
13.35 | -34.8396 | -22.3659
13.3561 | -34.84 | -22.3743
13.3622 | -34.8405 | -22.3829
13.3683 | -34.841 | -22.3912
13.3744 | -34.8415 | -22.3999
13.3805 | -34.8419 | -22.4083
13.3866 | -34.8424 | -22.4168
13.3927 | -34.8428 | -22.4256
13.3989 | -34.8433 | -22.4341
13.405 | -34.8438 | -22.4428
13.4111 | -34.8443 | -22.4515
13.4172 | -34.8447 | -22.46
13.4233 | -34.8452 | -22.4685
13.4294 | -34.8457 | -22.4772
13.4355 | -34.8461 | -22.4855
13.4416 | -34.8466 | -22.4939
13.4477 | -34.8472 | -22.5027
13.4538 | -34.8477 | -22.5113
13.4599 | -34.8482 | -22.52
13.466 | -34.8487 | -22.5285
13.4721 | -34.8491 | -22.5368
13.4782 | -34.8496 | -22.5454
13.4843 | -34.8501 | -22.5539
13.4904 | -34.8505 | -22.5623
13.4965 | -34.851 | -22.5709
13.5026 | -34.8515 | -22.5798
13.5087 | -34.852 | -22.588
13.5148 | -34.8524 | -22.5965
13.5209 | -34.8529 | -22.6055
13.527 | -34.8534 | -22.6141
13.5331 | -34.8538 | -22.6226
13.5393 | -34.8543 | -22.6309
13.5454 | -34.8548 | -22.6395
13.5515 | -34.8553 | -22.6484
13.5576 | -34.8558 | -22.657
13.5637 | -34.8562 | -22.6657
13.5698 | -34.8567 | -22.6744
13.5759 | -34.8572 | -22.6825
13.582 | -34.8577 | -22.691
13.5881 | -34.8582 | -22.6995
13.5942 | -34.8586 | -22.7081
13.6003 | -34.8591 | -22.7165
13.6064 | -34.8596 | -22.7253
13.6125 | -34.8601 | -22.7336
13.6186 | -34.8606 | -22.7419
13.6247 | -34.861 | -22.7504
13.6308 | -34.8615 | -22.7589
13.6369 | -34.862 | -22.7677
13.643 | -34.8624 | -22.7763
13.6491 | -34.8629 | -22.7849
13.6552 | -34.8634 | -22.7935
13.6613 | -34.8639 | -22.8022
13.6674 | -34.8643 | -22.8108
13.6735 | -34.8648 | -22.8192
13.6796 | -34.8653 | -22.8279
13.6858 | -34.8658 | -22.8363
13.6919 | -34.8663 | -22.8448
13.698 | -34.8667 | -22.8534
13.7041 | -34.8672 | -22.8622
13.7102 | -34.8677 | -22.8708
13.7163 | -34.8682 | -22.8792
13.7224 | -34.8687 | -22.8877
13.7285 | -34.8692 | -22.8963
13.7346 | -34.8696 | -22.9051
13.7407 | -34.8701 | -22.9135
13.7468 | -34.8706 | -22.922
13.7529 | -34.871 | -22.9304
13.759 | -34.8715 | -22.9389
13.7651 | -34.872 | -22.9474
13.7712 | -34.8725 | -22.956
13.7773 | -34.8729 | -22.9645
13.7834 | -34.8734 | -22.973
13.7895 | -34.8739 | -22.9814
13.7956 | -34.8744 | -22.9901
13.8017 | -34.8748 | -22.9985
13.8078 | -34.8753 | -23.007
13.8139 | -34.8757 | -23.0155
13.82 | -34.8762 | -23.0238
13.8262 | -34.8767 | -23.0324
13.8323 | -34.8772 | -23.0411
13.8384 | -34.8777 | -23.0496
13.8445 | -34.8782 | -23.0581
13.8506 | -34.8786 | -23.0667
13.8567 | -34.8791 | -23.0752
13.8628 | -34.8796 | -23.0837
13.8689 | -34.8801 | -23.0923
13.875 | -34.8806 | -23.1009
13.8811 | -34.881 | -23.1094
13.8872 | -34.8815 | -23.1178
13.8933 | -34.8819 | -23.126
13.8994 | -34.8824 | -23.1345
13.9055 | -34.8829 | -23.1432
13.9116 | -34.8833 | -23.1517
13.9177 | -34.8838 | -23.1601
13.9238 | -34.8842 | -23.1686
13.9299 | -34.8847 | -23.1768
13.936 | -34.8852 | -23.1853
13.9421 | -34.8856 | -23.1936
13.9482 | -34.8861 | -23.2021
13.9543 | -34.8865 | -23.2106
13.9604 | -34.887 | -23.219
13.9665 | -34.8875 | -23.2273
13.9727 | -34.888 | -23.2359
13.9788 | -34.8885 | -23.2443
13.9849 | -34.889 | -23.2528
13.991 | -34.8895 | -23.2614
13.9971 | -34.8899 | -23.2697
14.0032 | -34.8904 | -23.2782
14.0093 | -34.8909 | -23.2867
14.0154 | -34.8913 | -23.2951
14.0215 | -34.8918 | -23.3035
14.0276 | -34.8923 | -23.3121
14.0337 | -34.8928 | -23.3203
14.0398 | -34.8933 | -23.3286
14.0459 | -34.8937 | -23.337
14.052 | -34.8942 | -23.3456
14.0581 | -34.8947 | -23.354
14.0642 | -34.8952 | -23.3625
14.0703 | -34.8957 | -23.3709
14.0764 | -34.8961 | -23.3791
14.0825 | -34.8966 | -23.3877
14.0886 | -34.8971 | -23.3963
14.0947 | -34.8976 | -23.4047
14.1008 | -34.8981 | -23.4132
14.1069 | -34.8985 | -23.4213
14.1131 | -34.899 | -23.4297
14.1192 | -34.8995 | -23.4381
14.1253 | -34.8999 | -23.4464
14.1314 | -34.9004 | -23.4547
14.1375 | -34.9009 | -23.4633
14.1436 | -34.9014 | -23.4717
14.1497 | -34.9019 | -23.48
14.1558 | -34.9024 | -23.4886
14.1619 | -34.9029 | -23.497
14.168 | -34.9034 | -23.5055
14.1741 | -34.9039 | -23.5136
14.1802 | -34.9044 | -23.522
14.1863 | -34.9048 | -23.5303
14.1924 | -34.9053 | -23.5387
14.1985 | -34.9058 | -23.5471
14.2046 | -34.9063 | -23.5556
14.2107 | -34.9068 | -23.5638
14.2168 | -34.9073 | -23.5723
14.2229 | -34.9078 | -23.5806
14.229 | -34.9083 | -23.5891
14.2351 | -34.9087 | -23.5975
14.2412 | -34.9092 | -23.6059
14.2473 | -34.9097 | -23.6143
14.2534 | -34.9102 | -23.6225
14.2596 | -34.9106 | -23.6309
14.2657 | -34.9111 | -23.6393
14.2718 | -34.9116 | -23.6477
14.2779 | -34.9121 | -23.6562
14.284 | -34.9126 | -23.6645
14.2901 | -34.9131 | -23.6729
14.2962 | -34.9136 | -23.6813
14.3023 | -34.9141 | -23.6898
14.3084 | -34.9145 | -23.6982
14.3145 | -34.915 | -23.7065
14.3206 | -34.9155 | -23.715
14.3267 | -34.916 | -23.7234
14.3328 | -34.9165 | -23.7316
14.3389 | -34.917 | -23.74
14.345 | -34.9174 | -23.7483
14.3511 | -34.9179 | -23.7566
14.3572 | -34.9184 | -23.7649
14.3633 | -34.9189 | -23.7733
14.3694 | -34.9194 | -23.7817
14.3755 | -34.9199 | -23.79
14.3816 | -34.9204 | -23.7985
14.3877 | -34.9208 | -23.8068
14.3938 | -34.9213 | -23.8152
14.4 | -34.9218 | -23.8235
14.4061 | -34.9223 | -23.8318
14.4122 | -34.9228 | -23.8402
14.4183 | -34.9232 | -23.8484
14.4244 | -34.9237 | -23.8568
14.4305 | -34.9243 | -23.8651
14.4366 | -34.9248 | -23.8734
14.4427 | -34.9252 | -23.882
14.4488 | -34.9257 | -23.8906
14.4549 | -34.9262 | -23.8988
14.461 | -34.9267 | -23.9071
14.4671 | -34.9272 | -23.9155
14.4732 | -34.9277 | -23.9239
14.4793 | -34.9282 | -23.9323
14.4854 | -34.9287 | -23.9407
14.4915 | -34.9292 | -23.9487
14.4976 | -34.9297 | -23.957
14.5037 | -34.9302 | -23.9653
14.5098 | -34.9307 | -23.9736
14.5159 | -34.9311 | -23.9821
14.522 | -34.9316 | -23.9902
14.5281 | -34.9321 | -23.9984
14.5342 | -34.9326 | -24.007
14.5403 | -34.9331 | -24.0152
14.5465 | -34.9336 | -24.0235
14.5526 | -34.9341 | -24.0321
14.5587 | -34.9346 | -24.0405
14.5648 | -34.9351 | -24.0487
14.5709 | -34.9356 | -24.057
14.577 | -34.936 | -24.0653
14.5831 | -34.9365 | -24.0737
14.5892 | -34.937 | -24.082
14.5953 | -34.9375 | -24.0901
14.6014 | -34.938 | -24.0984
14.6075 | -34.9385 | -24.1065
14.6136 | -34.9389 | -24.115
14.6197 | -34.9394 | -24.1234
14.6258 | -34.9399 | -24.1317
14.6319 | -34.9404 | -24.14
14.638 | -34.9409 | -24.1483
14.6441 | -34.9414 | -24.1567
14.6502 | -34.9419 | -24.1652
14.6563 | -34.9424 | -24.1734
14.6624 | -34.9429 | -24.1819
14.6685 | -34.9434 | -24.1899
14.6746 | -34.9439 | -24.1982
14.6807 | -34.9444 | -24.2064
14.6869 | -34.9449 | -24.2146
14.693 | -34.9454 | -24.2229
14.6991 | -34.9459 | -24.2311
14.7052 | -34.9464 | -24.2393
14.7113 | -34.947 | -24.2478
14.7174 | -34.9474 | -24.2561
14.7235 | -34.948 | -24.2644
14.7296 | -34.9485 | -24.2725
14.7357 | -34.949 | -24.2806
14.7418 | -34.9495 | -24.2889
14.7479 | -34.95 | -24.2971
14.754 | -34.9505 | -24.3055
14.7601 | -34.951 | -24.3137
14.7662 | -34.9515 | -24.3218
14.7723 | -34.952 | -24.3299
14.7784 | -34.9525 | -24.3384
14.7845 | -34.953 | -24.3468
14.7906 | -34.9535 | -24.3551
14.7967 | -34.954 | -24.3634
14.8028 | -34.9545 | -24.3717
14.8089 | -34.955 | -24.3799
14.815 | -34.9556 | -24.3881
14.8211 | -34.956 | -24.3966
14.8272 | -34.9565 | -24.4048
14.8334 | -34.9571 | -24.4131
14.8395 | -34.9576 | -24.4211
14.8456 | -34.9581 | -24.4293
14.8517 | -34.9586 | -24.4376
14.8578 | -34.9591 | -24.4461
14.8639 | -34.9596 | -24.4545
14.87 | -34.9602 | -24.4626
14.8761 | -34.9607 | -24.4707
14.8822 | -34.9612 | -24.479
14.8883 | -34.9617 | -24.4872
14.8944 | -34.9622 | -24.4952
14.9005 | -34.9627 | -24.5035
14.9066 | -34.9632 | -24.5118
14.9127 | -34.9638 | -24.5201
14.9188 | -34.9642 | -24.5283
14.9249 | -34.9647 | -24.5368
14.931 | -34.9652 | -24.5451
14.9371 | -34.9657 | -24.5532
14.9432 | -34.9662 | -24.5615
14.9493 | -34.9668 | -24.5697
14.9554 | -34.9673 | -24.578
14.9615 | -34.9678 | -24.5861
14.9676 | -34.9683 | -24.5945
14.9738 | -34.9688 | -24.6027
14.9799 | -34.9693 | -24.6109
14.986 | -34.9699 | -24.6191
14.9921 | -34.9704 | -24.6273
14.9982 | -34.9709 | -24.6354
15.0043 | -34.9714 | -24.6436
15.0104 | -34.9719 | -24.6519
15.0165 | -34.9724 | -24.6603
15.0226 | -34.973 | -24.6683
15.0287 | -34.9735 | -24.6769
15.0348 | -34.974 | -24.6853
15.0409 | -34.9745 | -24.6934
15.047 | -34.975 | -24.7018
15.0531 | -34.9756 | -24.7098
15.0592 | -34.9761 | -24.7182
15.0653 | -34.9767 | -24.7266
15.0714 | -34.9771 | -24.7345
15.0775 | -34.9776 | -24.7426
15.0836 | -34.9781 | -24.7508
15.0897 | -34.9787 | -24.7588
15.0958 | -34.9792 | -24.767
15.1019 | -34.9798 | -24.7752
15.108 | -34.9803 | -24.7834
15.1141 | -34.9808 | -24.7916
15.1203 | -34.9813 | -24.7999
15.1264 | -34.9819 | -24.808
15.1325 | -34.9824 | -24.816
15.1386 | -34.9829 | -24.8241
15.1447 | -34.9834 | -24.8323
15.1508 | -34.984 | -24.8404
15.1569 | -34.9845 | -24.8488
15.163 | -34.9851 | -24.8567
15.1691 | -34.9856 | -24.8649
15.1752 | -34.9861 | -24.8731
15.1813 | -34.9867 | -24.8812
15.1874 | -34.9872 | -24.8893
15.1935 | -34.9877 | -24.8973
15.1996 | -34.9883 | -24.9053
15.2057 | -34.9888 | -24.9137
15.2118 | -34.9893 | -24.9217
15.2179 | -34.9899 | -24.9299
15.224 | -34.9904 | -24.9381
15.2301 | -34.991 | -24.9461
15.2362 | -34.9915 | -24.9545
15.2423 | -34.9921 | -24.9629
15.2484 | -34.9926 | -24.971
15.2545 | -34.9932 | -24.9792
15.2607 | -34.9937 | -24.9873
15.2668 | -34.9943 | -24.9952
15.2729 | -34.9948 | -25.0032
15.279 | -34.9954 | -25.0115
15.2851 | -34.9959 | -25.0195
15.2912 | -34.9964 | -25.0279
15.2973 | -34.997 | -25.036
15.3034 | -34.9975 | -25.044
15.3095 | -34.998 | -25.0522
15.3156 | -34.9986 | -25.0603
15.3217 | -34.9991 | -25.0683
15.3278 | -34.9997 | -25.0764
15.3339 | -35.0003 | -25.0843
15.34 | -35.0008 | -25.0923
15.3461 | -35.0013 | -25.1003
15.3522 | -35.0019 | -25.1084
15.3583 | -35.0023 | -25.1167
15.3644 | -35.0029 | -25.1246
15.3705 | -35.0035 | -25.1324
15.3766 | -35.004 | -25.1404
15.3827 | -35.0046 | -25.1484
15.3888 | -35.0052 | -25.1568
15.3949 | -35.0057 | -25.1647
15.401 | -35.0062 | -25.1729
15.4072 | -35.0068 | -25.1809
15.4133 | -35.0073 | -25.1893
15.4194 | -35.0078 | -25.1972
15.4255 | -35.0084 | -25.2053
15.4316 | -35.0089 | -25.2134
15.4377 | -35.0095 | -25.2214
15.4438 | -35.01 | -25.2293
15.4499 | -35.0106 | -25.2375
15.456 | -35.0112 | -25.2454
15.4621 | -35.0117 | -25.2537
15.4682 | -35.0123 | -25.262
15.4743 | -35.0128 | -25.2699
15.4804 | -35.0133 | -25.2781
15.4865 | -35.0139 | -25.286
15.4926 | -35.0144 | -25.2942
15.4987 | -35.015 | -25.3022
15.5048 | -35.0155 | -25.3103
15.5109 | -35.0161 | -25.3183
15.517 | -35.0167 | -25.3264
15.5231 | -35.0172 | -25.3344
15.5292 | -35.0177 | -25.3426
15.5353 | -35.0183 | -25.3509
15.5414 | -35.0188 | -25.3589
15.5476 | -35.0194 | -25.3666
15.5537 | -35.0199 | -25.3749
15.5598 | -35.0205 | -25.3829
15.5659 | -35.021 | -25.3909
15.572 | -35.0216 | -25.3989
15.5781 | -35.0221 | -25.4072
15.5842 | -35.0227 | -25.4151
15.5903 | -35.0233 | -25.4231
15.5964 | -35.0239 | -25.431
15.6025 | -35.0244 | -25.4391
15.6086 | -35.025 | -25.4473
15.6147 | -35.0255 | -25.4554
15.6208 | -35.0261 | -25.4632
15.6269 | -35.0266 | -25.4712
15.633 | -35.0271 | -25.479
15.6391 | -35.0277 | -25.4873
15.6452 | -35.0283 | -25.4951
15.6513 | -35.0289 | -25.503
15.6574 | -35.0294 | -25.511
15.6635 | -35.0299 | -25.5189
15.6696 | -35.0305 | -25.5271
15.6757 | -35.031 | -25.5352
15.6818 | -35.0316 | -25.5434
15.688 | -35.0321 | -25.5514
15.6941 | -35.0327 | -25.5595
15.7002 | -35.0332 | -25.5672
15.7063 | -35.0337 | -25.5752
15.7124 | -35.0343 | -25.5832
15.7185 | -35.0348 | -25.5913
15.7246 | -35.0354 | -25.5992
15.7307 | -35.036 | -25.607
15.7368 | -35.0365 | -25.6151
15.7429 | -35.0371 | -25.623
15.749 | -35.0377 | -25.631
15.7551 | -35.0382 | -25.639
15.7612 | -35.0388 | -25.6468
15.7673 | -35.0394 | -25.6547
15.7734 | -35.04 | -25.6629
15.7795 | -35.0405 | -25.671
15.7856 | -35.0411 | -25.6787
15.7917 | -35.0417 | -25.6866
15.7978 | -35.0423 | -25.6949
15.8039 | -35.0428 | -25.7029
15.81 | -35.0434 | -25.7111
15.8161 | -35.0439 | -25.7192
15.8222 | -35.0445 | -25.7273
15.8283 | -35.0451 | -25.7351
15.8345 | -35.0456 | -25.7432
15.8406 | -35.0462 | -25.7512
15.8467 | -35.0467 | -25.7591
15.8528 | -35.0473 | -25.7673
15.8589 | -35.0479 | -25.7754
15.865 | -35.0485 | -25.7831
15.8711 | -35.049 | -25.7913
15.8772 | -35.0496 | -25.7991
15.8833 | -35.0501 | -25.807
15.8894 | -35.0507 | -25.8151
15.8955 | -35.0512 | -25.8231
15.9016 | -35.0518 | -25.8308
15.9077 | -35.0524 | -25.8387
15.9138 | -35.0529 | -25.8468
15.9199 | -35.0534 | -25.8544
15.926 | -35.054 | -25.8622
15.9321 | -35.0546 | -25.8701
15.9382 | -35.0551 | -25.8776
15.9443 | -35.0557 | -25.8855
15.9504 | -35.0563 | -25.8933
15.9565 | -35.0568 | -25.9011
15.9626 | -35.0573 | -25.9092
15.9687 | -35.0579 | -25.9168
15.9749 | -35.0585 | -25.9247
15.981 | -35.059 | -25.9327
15.9871 | -35.0596 | -25.9407
15.9932 | -35.0602 | -25.9485
15.9993 | -35.0608 | -25.9565
16.0054 | -35.0614 | -25.9641
16.0115 | -35.0619 | -25.9721
16.0176 | -35.0625 | -25.98
16.0237 | -35.0631 | -25.9881
16.0298 | -35.0637 | -25.9959
16.0359 | -35.0642 | -26.0038
16.042 | -35.0648 | -26.0117
16.0481 | -35.0654 | -26.0198
16.0542 | -35.066 | -26.0277
16.0603 | -35.0665 | -26.0354
16.0664 | -35.0671 | -26.0433
16.0725 | -35.0676 | -26.0509
16.0786 | -35.0682 | -26.0589
16.0847 | -35.0688 | -26.067
16.0908 | -35.0694 | -26.0751
16.0969 | -35.0699 | -26.083
16.103 | -35.0705 | -26.0911
16.1091 | -35.0711 | -26.0989
16.1152 | -35.0716 | -26.1069
16.1214 | -35.0722 | -26.1149
16.1275 | -35.0727 | -26.1229
16.1336 | -35.0733 | -26.1308
16.1397 | -35.0739 | -26.1387
16.1458 | -35.0745 | -26.1464
16.1519 | -35.0751 | -26.1545
16.158 | -35.0757 | -26.1625
16.1641 | -35.0763 | -26.1704
16.1702 | -35.0768 | -26.1785
16.1763 | -35.0774 | -26.1864
16.1824 | -35.078 | -26.1942
16.1885 | -35.0785 | -26.202
16.1946 | -35.0791 | -26.2099
16.2007 | -35.0797 | -26.2179
16.2068 | -35.0802 | -26.2255
16.2129 | -35.0808 | -26.2334
16.219 | -35.0814 | -26.2409
16.2251 | -35.082 | -26.2487
16.2312 | -35.0826 | -26.2563
16.2373 | -35.0831 | -26.2642
16.2434 | -35.0838 | -26.2722
16.2495 | -35.0844 | -26.28
16.2556 | -35.085 | -26.2879
16.2618 | -35.0855 | -26.2958
16.2679 | -35.0861 | -26.3038
16.274 | -35.0866 | -26.3116
16.2801 | -35.0872 | -26.3196
16.2862 | -35.0878 | -26.3272
16.2923 | -35.0884 | -26.3351
16.2984 | -35.0889 | -26.3431
16.3045 | -35.0895 | -26.3511
16.3106 | -35.09 | -26.3587
16.3167 | -35.0907 | -26.3664
16.3228 | -35.0912 | -26.3739
16.3289 | -35.0918 | -26.3817
16.335 | -35.0924 | -26.3895
16.3411 | -35.093 | -26.3974
16.3472 | -35.0935 | -26.4055
16.3533 | -35.0941 | -26.4134
16.3594 | -35.0947 | -26.421
16.3655 | -35.0952 | -26.4288
16.3716 | -35.0958 | -26.4365
16.3777 | -35.0964 | -26.4442
16.3838 | -35.0969 | -26.452
16.3899 | -35.0975 | -26.4597
16.396 | -35.0981 | -26.4675
16.4021 | -35.0986 | -26.4751
16.4083 | -35.0992 | -26.4829
16.4144 | -35.0998 | -26.4907
16.4205 | -35.1004 | -26.4983
16.4266 | -35.1009 | -26.5059
16.4327 | -35.1015 | -26.5135
16.4388 | -35.102 | -26.5213
16.4449 | -35.1026 | -26.529
16.451 | -35.1032 | -26.5369
16.4571 | -35.1038 | -26.5446
16.4632 | -35.1043 | -26.5521
16.4693 | -35.1049 | -26.56
16.4754 | -35.1055 | -26.5677
16.4815 | -35.106 | -26.5756
16.4876 | -35.1066 | -26.5831
16.4937 | -35.1072 | -26.5908
16.4998 | -35.1077 | -26.5984
16.5059 | -35.1083 | -26.6061
16.512 | -35.1089 | -26.6139
16.5181 | -35.1095 | -26.622
16.5242 | -35.11 | -26.63
16.5303 | -35.1106 | -26.6376
16.5364 | -35.1112 | -26.6454
16.5425 | -35.1117 | -26.6532
16.5487 | -35.1123 | -26.6607
16.5548 | -35.1129 | -26.6688
16.5609 | -35.1135 | -26.6765
16.567 | -35.114 | -26.6841
16.5731 | -35.1146 | -26.6921
16.5792 | -35.1152 | -26.6998
16.5853 | -35.1158 | -26.7075
16.5914 | -35.1164 | -26.7153
16.5975 | -35.1169 | -26.7231
16.6036 | -35.1175 | -26.7307
16.6097 | -35.1181 | -26.7384
16.6158 | -35.1186 | -26.746
16.6219 | -35.1193 | -26.7539
16.628 | -35.1198 | -26.7616
16.6341 | -35.1204 | -26.7693
16.6402 | -35.121 | -26.7769
16.6463 | -35.1215 | -26.7846
16.6524 | -35.1221 | -26.7924
16.6585 | -35.1227 | -26.8001
16.6646 | -35.1233 | -26.8077
16.6707 | -35.1238 | -26.8154
16.6768 | -35.1244 | -26.8231
16.6829 | -35.125 | -26.8307
16.689 | -35.1256 | -26.8384
16.6952 | -35.1261 | -26.8462
16.7013 | -35.1267 | -26.8538
16.7074 | -35.1273 | -26.8612
16.7135 | -35.1279 | -26.8689
16.7196 | -35.1284 | -26.8764
16.7257 | -35.1291 | -26.8841
16.7318 | -35.1297 | -26.8917
16.7379 | -35.1302 | -26.8994
16.744 | -35.1308 | -26.907
16.7501 | -35.1314 | -26.9148
16.7562 | -35.1319 | -26.9225
16.7623 | -35.1325 | -26.9301
16.7684 | -35.133 | -26.938
16.7745 | -35.1336 | -26.9458
16.7806 | -35.1342 | -26.9532
16.7867 | -35.1347 | -26.9609
16.7928 | -35.1353 | -26.9686
16.7989 | -35.1359 | -26.9763
16.805 | -35.1364 | -26.9839
16.8111 | -35.137 | -26.9915
16.8172 | -35.1376 | -26.9993
16.8233 | -35.1381 | -27.007
16.8294 | -35.1387 | -27.0149
16.8356 | -35.1393 | -27.0226
16.8417 | -35.1399 | -27.0303
16.8478 | -35.1405 | -27.0378
16.8539 | -35.141 | -27.0455
16.86 | -35.1416 | -27.0532
16.8661 | -35.1422 | -27.0608
16.8722 | -35.1427 | -27.0686
16.8783 | -35.1433 | -27.0764
16.8844 | -35.1438 | -27.0839
16.8905 | -35.1444 | -27.0916
16.8966 | -35.145 | -27.0991
16.9027 | -35.1455 | -27.1068
16.9088 | -35.1461 | -27.1145
16.9149 | -35.1467 | -27.1224
16.921 | -35.1473 | -27.1298
16.9271 | -35.1478 | -27.1374
16.9332 | -35.1484 | -27.1451
16.9393 | -35.1489 | -27.1529
16.9454 | -35.1495 | -27.1606
16.9515 | -35.1501 | -27.1683
16.9576 | -35.1506 | -27.1758
16.9637 | -35.1512 | -27.1835
16.9698 | -35.1518 | -27.191
16.9759 | -35.1523 | -27.1986
16.9821 | -35.1529 | -27.2065
16.9882 | -35.1535 | -27.2139
16.9943 | -35.1541 | -27.2214
17.0004 | -35.1546 | -27.2291
17.0065 | -35.1552 | -27.2369
17.0126 | -35.1558 | -27.2445
17.0187 | -35.1563 | -27.2523
17.0248 | -35.1569 | -27.2598
17.0309 | -35.1575 | -27.2673
17.037 | -35.158 | -27.2751
17.0431 | -35.1586 | -27.2827
17.0492 | -35.1592 | -27.2903
17.0553 | -35.1597 | -27.2978
17.0614 | -35.1603 | -27.3054
17.0675 | -35.1609 | -27.3131
17.0736 | -35.1614 | -27.3207
17.0797 | -35.162 | -27.3283
17.0858 | -35.1626 | -27.3362
17.0919 | -35.1631 | -27.3438
17.098 | -35.1637 | -27.3512
17.1041 | -35.1643 | -27.3587
17.1102 | -35.1648 | -27.3661
17.1163 | -35.1654 | -27.3738
17.1225 | -35.166 | -27.3812
17.1286 | -35.1665 | -27.3886
17.1347 | -35.1671 | -27.3963
17.1408 | -35.1677 | -27.4039
17.1469 | -35.1683 | -27.4114
17.153 | -35.1688 | -27.4191
17.1591 | -35.1694 | -27.4266
17.1652 | -35.17 | -27.434
17.1713 | -35.1705 | -27.4417
17.1774 | -35.1711 | -27.4493
17.1835 | -35.1717 | -27.4569
17.1896 | -35.1722 | -27.4645
17.1957 | -35.1728 | -27.472
17.2018 | -35.1733 | -27.4793
17.2079 | -35.1739 | -27.4868
17.214 | -35.1744 | -27.4946
17.2201 | -35.175 | -27.5022
17.2262 | -35.1756 | -27.5098
17.2323 | -35.1762 | -27.5175
17.2384 | -35.1768 | -27.5249
17.2445 | -35.1773 | -27.5325
17.2506 | -35.1779 | -27.5402
17.2567 | -35.1784 | -27.5478
17.2628 | -35.179 | -27.5554
17.269 | -35.1795 | -27.5629
17.2751 | -35.1801 | -27.5704
17.2812 | -35.1806 | -27.5783
17.2873 | -35.1812 | -27.586
17.2934 | -35.1818 | -27.5936
17.2995 | -35.1824 | -27.6013
17.3056 | -35.1829 | -27.6087
17.3117 | -35.1835 | -27.6164
17.3178 | -35.184 | -27.6238
17.3239 | -35.1846 | -27.6314
17.33 | -35.1851 | -27.639
17.3361 | -35.1857 | -27.6464
17.3422 | -35.1862 | -27.6537
17.3483 | -35.1868 | -27.6613
17.3544 | -35.1874 | -27.6689
17.3605 | -35.1879 | -27.6764
17.3666 | -35.1885 | -27.6839
17.3727 | -35.1891 | -27.6913
17.3788 | -35.1896 | -27.6988
17.3849 | -35.1902 | -27.7063
17.391 | -35.1907 | -27.714
17.3971 | -35.1913 | -27.7215
17.4032 | -35.1919 | -27.7291
17.4094 | -35.1924 | -27.7365
17.4155 | -35.193 | -27.744
17.4216 | -35.1935 | -27.7515
17.4277 | -35.1941 | -27.759
17.4338 | -35.1947 | -27.7665
17.4399 | -35.1952 | -27.7741
17.446 | -35.1958 | -27.7813
17.4521 | -35.1963 | -27.7888
17.4582 | -35.1969 | -27.7962
17.4643 | -35.1974 | -27.8038
17.4704 | -35.198 | -27.8112
17.4765 | -35.1986 | -27.8188
17.4826 | -35.1991 | -27.8263
17.4887 | -35.1997 | -27.8338
17.4948 | -35.2002 | -27.8414
17.5009 | -35.2008 | -27.8489
17.507 | -35.2013 | -27.8564
17.5131 | -35.2018 | -27.864
17.5192 | -35.2024 | -27.8713
17.5253 | -35.2029 | -27.8787
17.5314 | -35.2035 | -27.8863
17.5375 | -35.204 | -27.8938
17.5436 | -35.2046 | -27.9015
17.5497 | -35.2051 | -27.9089
17.5559 | -35.2057 | -27.9161
17.562 | -35.2062 | -27.9234
17.5681 | -35.2068 | -27.931
17.5742 | -35.2074 | -27.9387
17.5803 | -35.2079 | -27.9461
17.5864 | -35.2084 | -27.9535
17.5925 | -35.209 | -27.961
17.5986 | -35.2096 | -27.9684
17.6047 | -35.2102 | -27.9758
17.6108 | -35.2107 | -27.9833
17.6169 | -35.2113 | -27.9909
17.623 | -35.2118 | -27.9982
17.6291 | -35.2124 | -28.0054
17.6352 | -35.2129 | -28.0128
17.6413 | -35.2135 | -28.0203
17.6474 | -35.214 | -28.0281
17.6535 | -35.2146 | -28.0356
17.6596 | -35.2152 | -28.0428
17.6657 | -35.2157 | -28.0504
17.6718 | -35.2163 | -28.0579
17.6779 | -35.2168 | -28.0653
17.684 | -35.2174 | -28.0728
17.6901 | -35.218 | -28.0804
17.6963 | -35.2185 | -28.0876
17.7024 | -35.2191 | -28.095
17.7085 | -35.2196 | -28.1024
17.7146 | -35.2202 | -28.1099
17.7207 | -35.2207 | -28.1173
17.7268 | -35.2213 | -28.1246
17.7329 | -35.2218 | -28.132
17.739 | -35.2224 | -28.1396
17.7451 | -35.223 | -28.1469
17.7512 | -35.2235 | -28.1544
17.7573 | -35.2241 | -28.162
17.7634 | -35.2246 | -28.1692
17.7695 | -35.2252 | -28.1767
17.7756 | -35.2258 | -28.1843
17.7817 | -35.2263 | -28.1917
17.7878 | -35.2269 | -28.1993
17.7939 | -35.2275 | -28.2069
17.8 | -35.228 | -28.2141
17.8061 | -35.2286 | -28.2218
17.8122 | -35.2291 | -28.2293
17.8183 | -35.2297 | -28.2368
17.8244 | -35.2302 | -28.2443
17.8305 | -35.2308 | -28.2517
17.8366 | -35.2313 | -28.2589
17.8428 | -35.2319 | -28.2664
17.8489 | -35.2324 | -28.2737
17.855 | -35.233 | -28.281
17.8611 | -35.2335 | -28.2884
17.8672 | -35.2341 | -28.2958
17.8733 | -35.2347 | -28.3033
17.8794 | -35.2353 | -28.3107
17.8855 | -35.2358 | -28.3183
17.8916 | -35.2363 | -28.3259
17.8977 | -35.2369 | -28.3333
17.9038 | -35.2375 | -28.3404
17.9099 | -35.238 | -28.3478
17.916 | -35.2385 | -28.3551
17.9221 | -35.2391 | -28.3626
17.9282 | -35.2396 | -28.3699
17.9343 | -35.2402 | -28.3772
17.9404 | -35.2407 | -28.3846
17.9465 | -35.2412 | -28.3919
17.9526 | -35.2418 | -28.3994
17.9587 | -35.2423 | -28.4069
17.9648 | -35.2429 | -28.4142
17.9709 | -35.2434 | -28.4216
17.977 | -35.244 | -28.4289
17.9832 | -35.2445 | -28.4364
17.9893 | -35.2451 | -28.444
17.9954 | -35.2456 | -28.4516
18.0015 | -35.2461 | -28.459
18.0076 | -35.2467 | -28.4663
18.0137 | -35.2472 | -28.4739
18.0198 | -35.2477 | -28.4812
18.0259 | -35.2483 | -28.4889
18.032 | -35.2488 | -28.4964
18.0381 | -35.2494 | -28.5039
18.0442 | -35.2499 | -28.5111
18.0503 | -35.2504 | -28.5185
18.0564 | -35.251 | -28.5261
18.0625 | -35.2515 | -28.5335
18.0686 | -35.2521 | -28.541
18.0747 | -35.2526 | -28.5487
18.0808 | -35.2531 | -28.5559
18.0869 | -35.2537 | -28.5633
18.093 | -35.2542 | -28.5708
18.0991 | -35.2547 | -28.5782
18.1052 | -35.2552 | -28.5858
18.1113 | -35.2558 | -28.5933
18.1174 | -35.2563 | -28.6006
18.1236 | -35.2569 | -28.6083
18.1297 | -35.2574 | -28.6158
18.1358 | -35.258 | -28.623
18.1419 | -35.2585 | -28.6305
18.148 | -35.2591 | -28.6379
18.1541 | -35.2596 | -28.645
18.1602 | -35.2602 | -28.6526
18.1663 | -35.2607 | -28.6597
18.1724 | -35.2613 | -28.667
18.1785 | -35.2618 | -28.6745
18.1846 | -35.2623 | -28.6818
18.1907 | -35.2629 | -28.6891
18.1968 | -35.2634 | -28.6965
18.2029 | -35.2639 | -28.7038
18.209 | -35.2645 | -28.7112
18.2151 | -35.265 | -28.7186
18.2212 | -35.2655 | -28.7259
18.2273 | -35.266 | -28.7332
18.2334 | -35.2666 | -28.7405
18.2395 | -35.2671 | -28.7478
18.2456 | -35.2676 | -28.7554
18.2517 | -35.2681 | -28.7626
18.2578 | -35.2687 | -28.7699
18.2639 | -35.2692 | -28.7775
18.2701 | -35.2698 | -28.7849
18.2762 | -35.2704 | -28.792
18.2823 | -35.2709 | -28.7995
18.2884 | -35.2714 | -28.807
18.2945 | -35.2719 | -28.8141
18.3006 | -35.2725 | -28.8215
18.3067 | -35.273 | -28.8288
18.3128 | -35.2736 | -28.8363
18.3189 | -35.2741 | -28.8436
18.325 | -35.2746 | -28.8508
18.3311 | -35.2751 | -28.8581
18.3372 | -35.2757 | -28.8654
18.3433 | -35.2763 | -28.873
18.3494 | -35.2768 | -28.8805
18.3555 | -35.2773 | -28.8879
18.3616 | -35.2779 | -28.8951
18.3677 | -35.2784 | -28.9026
18.3738 | -35.279 | -28.9099
18.3799 | -35.2795 | -28.9172
18.386 | -35.28 | -28.9245
18.3921 | -35.2805 | -28.9321
18.3982 | -35.2811 | -28.9392
18.4043 | -35.2816 | -28.9465
18.4105 | -35.2821 | -28.9539
18.4166 | -35.2826 | -28.9613
18.4227 | -35.2831 | -28.9687
18.4288 | -35.2837 | -28.976
18.4349 | -35.2842 | -28.9832
18.441 | -35.2847 | -28.9905
18.4471 | -35.2853 | -28.9978
18.4532 | -35.2858 | -29.0052
18.4593 | -35.2864 | -29.0124
18.4654 | -35.2869 | -29.0196
18.4715 | -35.2874 | -29.0273
18.4776 | -35.2879 | -29.0349
18.4837 | -35.2885 | -29.0422
18.4898 | -35.289 | -29.0494
18.4959 | -35.2895 | -29.0569
18.502 | -35.29 | -29.0644
18.5081 | -35.2906 | -29.0718
18.5142 | -35.2911 | -29.0792
18.5203 | -35.2916 | -29.0866
18.5264 | -35.2922 | -29.0939
18.5325 | -35.2927 | -29.1012
18.5386 | -35.2932 | -29.1086
18.5447 | -35.2937 | -29.1159
18.5508 | -35.2943 | -29.1232
18.557 | -35.2948 | -29.1308
18.5631 | -35.2954 | -29.1383
18.5692 | -35.2959 | -29.1455
18.5753 | -35.2964 | -29.1527
18.5814 | -35.297 | -29.1598
18.5875 | -35.2975 | -29.1672
18.5936 | -35.298 | -29.1746
18.5997 | -35.2986 | -29.182
18.6058 | -35.2991 | -29.1891
18.6119 | -35.2996 | -29.1966
18.618 | -35.3001 | -29.2039
18.6241 | -35.3007 | -29.2114
18.6302 | -35.3012 | -29.2187
18.6363 | -35.3017 | -29.2261
18.6424 | -35.3023 | -29.2335
18.6485 | -35.3028 | -29.2409
18.6546 | -35.3034 | -29.2484
18.6607 | -35.3039 | -29.2557
18.6668 | -35.3044 | -29.2628
18.6729 | -35.3049 | -29.2702
18.679 | -35.3055 | -29.2775
18.6851 | -35.306 | -29.2847
18.6912 | -35.3065 | -29.2921
18.6974 | -35.3071 | -29.2994
18.7035 | -35.3076 | -29.3067
18.7096 | -35.3081 | -29.3141
18.7157 | -35.3087 | -29.3213
18.7218 | -35.3092 | -29.3286
18.7279 | -35.3098 | -29.3361
18.734 | -35.3103 | -29.3435
18.7401 | -35.3109 | -29.3506
18.7462 | -35.3114 | -29.3577
18.7523 | -35.3119 | -29.3651
18.7584 | -35.3124 | -29.3724
18.7645 | -35.313 | -29.3798
18.7706 | -35.3135 | -29.3872
18.7767 | -35.314 | -29.3942
18.7828 | -35.3146 | -29.4015
18.7889 | -35.3151 | -29.4089
18.795 | -35.3156 | -29.4166
18.8011 | -35.3162 | -29.4241
18.8072 | -35.3167 | -29.4315
18.8133 | -35.3172 | -29.439
18.8194 | -35.3177 | -29.446
18.8255 | -35.3182 | -29.4533
18.8316 | -35.3187 | -29.4605
18.8377 | -35.3192 | -29.4676
18.8439 | -35.3198 | -29.475
18.85 | -35.3203 | -29.4821
18.8561 | -35.3208 | -29.4894
18.8622 | -35.3212 | -29.497
18.8683 | -35.3218 | -29.5039
18.8744 | -35.3223 | -29.5111
18.8805 | -35.3228 | -29.5185
18.8866 | -35.3233 | -29.5259
18.8927 | -35.3238 | -29.533
18.8988 | -35.3244 | -29.5404
18.9049 | -35.3248 | -29.5475
18.911 | -35.3254 | -29.5546
18.9171 | -35.3259 | -29.5619
18.9232 | -35.3264 | -29.5689
18.9293 | -35.327 | -29.576
18.9354 | -35.3275 | -29.5835
18.9415 | -35.328 | -29.5906
18.9476 | -35.3285 | -29.5978
18.9537 | -35.329 | -29.605
18.9598 | -35.3295 | -29.6119
18.9659 | -35.33 | -29.6192
18.972 | -35.3306 | -29.6264
18.9781 | -35.3311 | -29.6339
18.9843 | -35.3316 | -29.6411
18.9904 | -35.3321 | -29.6483
18.9965 | -35.3327 | -29.6553
19.0026 | -35.3332 | -29.6629
19.0087 | -35.3338 | -29.67
19.0148 | -35.3343 | -29.6774
19.0209 | -35.3349 | -29.6847
19.027 | -35.3354 | -29.6921
19.0331 | -35.3359 | -29.6991
19.0392 | -35.3365 | -29.7064
19.0453 | -35.337 | -29.7136
19.0514 | -35.3376 | -29.7207
19.0575 | -35.3381 | -29.7279
19.0636 | -35.3386 | -29.7347
19.0697 | -35.3392 | -29.7419
19.0758 | -35.3397 | -29.7491
19.0819 | -35.3402 | -29.7563
19.088 | -35.3407 | -29.7637
19.0941 | -35.3412 | -29.771
19.1002 | -35.3417 | -29.778
19.1063 | -35.3422 | -29.7854
19.1124 | -35.3427 | -29.7924
19.1185 | -35.3432 | -29.7995
19.1246 | -35.3437 | -29.8068
19.1308 | -35.3442 | -29.814
19.1369 | -35.3448 | -29.8209
19.143 | -35.3453 | -29.8282
19.1491 | -35.3458 | -29.8355
19.1552 | -35.3464 | -29.8428
19.1613 | -35.3469 | -29.8501
19.1674 | -35.3474 | -29.8573
19.1735 | -35.3479 | -29.8644
19.1796 | -35.3485 | -29.8717
19.1857 | -35.349 | -29.8787
19.1918 | -35.3495 | -29.886
19.1979 | -35.35 | -29.8932
19.204 | -35.3506 | -29.9003
19.2101 | -35.3511 | -29.9076
19.2162 | -35.3516 | -29.9146
19.2223 | -35.3522 | -29.9218
19.2284 | -35.3527 | -29.9292
19.2345 | -35.3532 | -29.9365
19.2406 | -35.3537 | -29.9435
19.2467 | -35.3542 | -29.9508
19.2528 | -35.3547 | -29.958
19.2589 | -35.3552 | -29.965
19.265 | -35.3558 | -29.9723
19.2712 | -35.3563 | -29.9793
19.2773 | -35.3568 | -29.9865
19.2834 | -35.3574 | -29.9936
19.2895 | -35.3579 | -30.0009
19.2956 | -35.3584 | -30.0082
19.3017 | -35.3589 | -30.0153
19.3078 | -35.3594 | -30.0229
19.3139 | -35.3599 | -30.0304
19.32 | -35.3605 | -30.0375
19.3261 | -35.3609 | -30.0448
19.3322 | -35.3614 | -30.052
19.3383 | -35.362 | -30.0592
19.3444 | -35.3625 | -30.0662
19.3505 | -35.363 | -30.0735
19.3566 | -35.3635 | -30.0808
19.3627 | -35.364 | -30.0882
19.3688 | -35.3645 | -30.0953
19.3749 | -35.365 | -30.1026
19.381 | -35.3656 | -30.1093
19.3871 | -35.3661 | -30.1165
19.3932 | -35.3666 | -30.1238
19.3993 | -35.3672 | -30.1308
19.4054 | -35.3677 | -30.138
19.4115 | -35.3682 | -30.1454
19.4177 | -35.3687 | -30.1522
19.4238 | -35.3692 | -30.1591
19.4299 | -35.3697 | -30.1664
19.436 | -35.3703 | -30.1736
19.4421 | -35.3708 | -30.1807
19.4482 | -35.3713 | -30.188
19.4543 | -35.3718 | -30.1947
19.4604 | -35.3723 | -30.2019
19.4665 | -35.3728 | -30.209
19.4726 | -35.3733 | -30.216
19.4787 | -35.3738 | -30.223
19.4848 | -35.3743 | -30.2304
19.4909 | -35.3748 | -30.2373
19.497 | -35.3753 | -30.2445
19.5031 | -35.3759 | -30.2518
19.5092 | -35.3764 | -30.2588
19.5153 | -35.3769 | -30.2659
19.5214 | -35.3774 | -30.2728
19.5275 | -35.3779 | -30.28
19.5336 | -35.3784 | -30.2871
19.5397 | -35.379 | -30.2942
19.5458 | -35.3795 | -30.3014
19.5519 | -35.38 | -30.3084
19.5581 | -35.3805 | -30.3154
19.5642 | -35.381 | -30.3225
19.5703 | -35.3815 | -30.3296
19.5764 | -35.382 | -30.3367
19.5825 | -35.3825 | -30.3437
19.5886 | -35.3831 | -30.351
19.5947 | -35.3836 | -30.3581
19.6008 | -35.3841 | -30.3651
19.6069 | -35.3846 | -30.3722
19.613 | -35.3851 | -30.3795
19.6191 | -35.3856 | -30.3865
19.6252 | -35.3861 | -30.3938
19.6313 | -35.3866 | -30.4006
19.6374 | -35.3871 | -30.408
19.6435 | -35.3876 | -30.415
19.6496 | -35.3881 | -30.4219
19.6557 | -35.3886 | -30.4289
19.6618 | -35.3892 | -30.4356
19.6679 | -35.3897 | -30.4426
19.674 | -35.3902 | -30.4498
19.6801 | -35.3907 | -30.4567
19.6862 | -35.3912 | -30.4637
19.6923 | -35.3917 | -30.4707
19.6984 | -35.3922 | -30.4778
19.7046 | -35.3927 | -30.485
19.7107 | -35.3932 | -30.4921
19.7168 | -35.3937 | -30.499
19.7229 | -35.3942 | -30.5061
19.729 | -35.3947 | -30.5133
19.7351 | -35.3952 | -30.5201
19.7412 | -35.3958 | -30.5273
19.7473 | -35.3963 | -30.5345
19.7534 | -35.3968 | -30.5416
19.7595 | -35.3973 | -30.5487
19.7656 | -35.3978 | -30.5558
19.7717 | -35.3983 | -30.5626
19.7778 | -35.3988 | -30.5697
19.7839 | -35.3993 | -30.5768
19.79 | -35.3998 | -30.5839
19.7961 | -35.4003 | -30.5911
19.8022 | -35.4008 | -30.5979
19.8083 | -35.4014 | -30.6051
19.8144 | -35.4019 | -30.6122
19.8205 | -35.4024 | -30.6191
19.8266 | -35.4029 | -30.6262
19.8327 | -35.4034 | -30.6332
19.8388 | -35.4039 | -30.6402
19.845 | -35.4044 | -30.6473
19.8511 | -35.4049 | -30.6545
19.8572 | -35.4054 | -30.6615
19.8633 | -35.4059 | -30.6685
19.8694 | -35.4064 | -30.6758
19.8755 | -35.4069 | -30.6825
19.8816 | -35.4074 | -30.6897
19.8877 | -35.4079 | -30.6966
19.8938 | -35.4084 | -30.7037
19.8999 | -35.4088 | -30.7107
19.906 | -35.4093 | -30.718
19.9121 | -35.4098 | -30.7249
19.9182 | -35.4104 | -30.7321
19.9243 | -35.4109 | -30.7391
19.9304 | -35.4113 | -30.7461
19.9365 | -35.4118 | -30.7532
19.9426 | -35.4123 | -30.76
19.9487 | -35.4128 | -30.7672
19.9548 | -35.4133 | -30.774
19.9609 | -35.4138 | -30.7811
19.967 | -35.4143 | -30.788
19.9731 | -35.4148 | -30.7952
19.9792 | -35.4153 | -30.8021
19.9853 | -35.4158 | -30.809
19.9915 | -35.4163 | -30.8159
19.9976 | -35.4168 | -30.8231
20.0037 | -35.4173 | -30.8301
20.0098 | -35.4178 | -30.8371
20.0159 | -35.4184 | -30.8441
20.022 | -35.4189 | -30.8512
20.0281 | -35.4194 | -30.8582
20.0342 | -35.4199 | -30.8651
20.0403 | -35.4204 | -30.8724
20.0464 | -35.4209 | -30.8797
20.0525 | -35.4214 | -30.8866
20.0586 | -35.422 | -30.8937
20.0647 | -35.4224 | -30.9007
20.0708 | -35.4229 | -30.9081
20.0769 | -35.4235 | -30.9152
20.083 | -35.424 | -30.9219
20.0891 | -35.4245 | -30.9288
20.0952 | -35.425 | -30.9356
20.1013 | -35.4255 | -30.9428
20.1074 | -35.426 | -30.9499
20.1135 | -35.4265 | -30.957
20.1196 | -35.427 | -30.9637
20.1257 | -35.4275 | -30.9705
20.1319 | -35.428 | -30.9775
20.138 | -35.4285 | -30.9845
20.1441 | -35.429 | -30.9916
20.1502 | -35.4295 | -30.9986
20.1563 | -35.43 | -31.0055
20.1624 | -35.4306 | -31.0125
20.1685 | -35.4311 | -31.0196
20.1746 | -35.4315 | -31.0265
20.1807 | -35.432 | -31.0335
20.1868 | -35.4325 | -31.0405
20.1929 | -35.433 | -31.0472
20.199 | -35.4336 | -31.0542
20.2051 | -35.4341 | -31.0609
20.2112 | -35.4346 | -31.0679
20.2173 | -35.4351 | -31.0749
20.2234 | -35.4356 | -31.0819
20.2295 | -35.4361 | -31.0889
20.2356 | -35.4367 | -31.0958
20.2417 | -35.4372 | -31.1027
20.2478 | -35.4377 | -31.1095
20.2539 | -35.4382 | -31.1165
20.26 | -35.4387 | -31.1234
20.2661 | -35.4393 | -31.1306
20.2723 | -35.4398 | -31.1375
20.2784 | -35.4403 | -31.1443
20.2845 | -35.4408 | -31.1512
20.2906 | -35.4413 | -31.158
20.2967 | -35.4418 | -31.1649
20.3028 | -35.4423 | -31.1717
20.3089 | -35.4428 | -31.1787
20.315 | -35.4433 | -31.1858
20.3211 | -35.4438 | -31.1926
20.3272 | -35.4443 | -31.1997
20.3333 | -35.4448 | -31.2061
20.3394 | -35.4453 | -31.2129
20.3455 | -35.4458 | -31.2198
20.3516 | -35.4463 | -31.2265
20.3577 | -35.4468 | -31.2334
20.3638 | -35.4473 | -31.2404
20.3699 | -35.4478 | -31.2471
20.376 | -35.4483 | -31.2539
20.3821 | -35.4488 | -31.2609
20.3882 | -35.4493 | -31.2676
20.3943 | -35.4498 | -31.2744
20.4004 | -35.4503 | -31.2812
20.4065 | -35.4508 | -31.288
20.4126 | -35.4513 | -31.295
20.4188 | -35.4518 | -31.3021
20.4249 | -35.4523 | -31.309
20.431 | -35.4528 | -31.3157
20.4371 | -35.4532 | -31.3225
20.4432 | -35.4537 | -31.3293
20.4493 | -35.4542 | -31.3364
20.4554 | -35.4547 | -31.3432
20.4615 | -35.4552 | -31.3503
20.4676 | -35.4557 | -31.3572
20.4737 | -35.4562 | -31.3639
20.4798 | -35.4567 | -31.3708
20.4859 | -35.4572 | -31.378
20.492 | -35.4577 | -31.3846
20.4981 | -35.4582 | -31.3915
20.5042 | -35.4587 | -31.3983
20.5103 | -35.4592 | -31.405
20.5164 | -35.4597 | -31.4119
20.5225 | -35.4602 | -31.4187
20.5286 | -35.4607 | -31.4256
20.5347 | -35.4612 | -31.4323
20.5408 | -35.4618 | -31.439
20.5469 | -35.4623 | -31.446
20.553 | -35.4628 | -31.4527
20.5592 | -35.4633 | -31.4595
20.5653 | -35.4638 | -31.4664
20.5714 | -35.4644 | -31.4731
20.5775 | -35.4649 | -31.4799
20.5836 | -35.4654 | -31.4867
20.5897 | -35.4659 | -31.4934
20.5958 | -35.4664 | -31.5004
20.6019 | -35.4669 | -31.5071
20.608 | -35.4673 | -31.514
20.6141 | -35.4679 | -31.5206
20.6202 | -35.4684 | -31.5274
20.6263 | -35.4689 | -31.5341
20.6324 | -35.4694 | -31.5413
20.6385 | -35.4699 | -31.548
20.6446 | -35.4705 | -31.5549
20.6507 | -35.471 | -31.5616
20.6568 | -35.4715 | -31.5682
20.6629 | -35.472 | -31.5748
20.669 | -35.4725 | -31.5817
20.6751 | -35.473 | -31.5886
20.6812 | -35.4735 | -31.5949
20.6873 | -35.474 | -31.6018
20.6934 | -35.4745 | -31.6085
20.6995 | -35.475 | -31.6152
20.7057 | -35.4756 | -31.6221
20.7118 | -35.4761 | -31.6289
20.7179 | -35.4766 | -31.6351
20.724 | -35.4772 | -31.642
20.7301 | -35.4777 | -31.6489
20.7362 | -35.4782 | -31.6557
20.7423 | -35.4787 | -31.6625
20.7484 | -35.4792 | -31.6694
20.7545 | -35.4797 | -31.6761
20.7606 | -35.4802 | -31.6831
20.7667 | -35.4807 | -31.6897
20.7728 | -35.4812 | -31.6962
20.7789 | -35.4817 | -31.7032
20.785 | -35.4822 | -31.7099
20.7911 | -35.4827 | -31.7164
20.7972 | -35.4832 | -31.7231
20.8033 | -35.4837 | -31.7299
20.8094 | -35.4842 | -31.7365
20.8155 | -35.4848 | -31.7435
20.8216 | -35.4853 | -31.7504
20.8277 | -35.4858 | -31.7574
20.8338 | -35.4863 | -31.7645
20.8399 | -35.4868 | -31.7713
20.8461 | -35.4873 | -31.7782
20.8522 | -35.4878 | -31.7848
20.8583 | -35.4884 | -31.7913
20.8644 | -35.4889 | -31.7981
20.8705 | -35.4894 | -31.805
20.8766 | -35.4899 | -31.8119
20.8827 | -35.4904 | -31.8188
20.8888 | -35.4909 | -31.8256
20.8949 | -35.4914 | -31.8322
20.901 | -35.4919 | -31.839
20.9071 | -35.4924 | -31.8458
20.9132 | -35.493 | -31.8526
20.9193 | -35.4935 | -31.8595
20.9254 | -35.494 | -31.8662
20.9315 | -35.4945 | -31.8729
20.9376 | -35.495 | -31.8797
20.9437 | -35.4955 | -31.8866
20.9498 | -35.496 | -31.8934
20.9559 | -35.4965 | -31.9001
20.962 | -35.497 | -31.9068
20.9681 | -35.4976 | -31.9136
20.9742 | -35.4981 | -31.9202
20.9803 | -35.4986 | -31.9268
20.9864 | -35.4991 | -31.9336
20.9926 | -35.4996 | -31.9404
20.9987 | -35.5001 | -31.9469
21.0048 | -35.5006 | -31.9534
21.0109 | -35.5012 | -31.9603
21.017 | -35.5017 | -31.9668
21.0231 | -35.5022 | -31.9738
21.0292 | -35.5027 | -31.9805
21.0353 | -35.5033 | -31.9866
21.0414 | -35.5038 | -31.9933
21.0475 | -35.5043 | -32
21.0536 | -35.5048 | -32.0067
21.0597 | -35.5053 | -32.0134
21.0658 | -35.5058 | -32.02
21.0719 | -35.5063 | -32.0262
21.078 | -35.5068 | -32.0331
21.0841 | -35.5072 | -32.0399
21.0902 | -35.5077 | -32.0462
21.0963 | -35.5082 | -32.0531
21.1024 | -35.5087 | -32.0598
21.1085 | -35.5092 | -32.0665
21.1146 | -35.5097 | -32.0733
21.1207 | -35.5102 | -32.0799
21.1268 | -35.5107 | -32.0865
21.133 | -35.5112 | -32.0934
21.1391 | -35.5118 | -32.1001
21.1452 | -35.5123 | -32.1067
21.1513 | -35.5128 | -32.1132
21.1574 | -35.5133 | -32.1198
21.1635 | -35.5138 | -32.1263
21.1696 | -35.5144 | -32.1328
21.1757 | -35.5149 | -32.1393
21.1818 | -35.5154 | -32.146
21.1879 | -35.516 | -32.1529
21.194 | -35.5164 | -32.1598
21.2001 | -35.517 | -32.1665
21.2062 | -35.5175 | -32.1732
21.2123 | -35.518 | -32.1798
21.2184 | -35.5185 | -32.1863
21.2245 | -35.5191 | -32.193
21.2306 | -35.5196 | -32.1994
21.2367 | -35.52 | -32.2059
21.2428 | -35.5206 | -32.2123
21.2489 | -35.5211 | -32.2188
21.255 | -35.5216 | -32.2257
21.2611 | -35.5221 | -32.2327
21.2672 | -35.5226 | -32.2396
21.2733 | -35.5231 | -32.246
21.2795 | -35.5236 | -32.2522
21.2856 | -35.5241 | -32.2589
21.2917 | -35.5246 | -32.2656
21.2978 | -35.5251 | -32.2726
21.3039 | -35.5256 | -32.2792
21.31 | -35.5261 | -32.286
21.3161 | -35.5266 | -32.2923
21.3222 | -35.5271 | -32.2991
21.3283 | -35.5276 | -32.3057
21.3344 | -35.5281 | -32.3123
21.3405 | -35.5286 | -32.319
21.3466 | -35.5291 | -32.3258
21.3527 | -35.5296 | -32.3321
21.3588 | -35.5302 | -32.3388
21.3649 | -35.5307 | -32.3457
21.371 | -35.5312 | -32.3523
21.3771 | -35.5317 | -32.3589
21.3832 | -35.5323 | -32.3654
21.3893 | -35.5328 | -32.372
21.3954 | -35.5334 | -32.3787
21.4015 | -35.5339 | -32.3849
21.4076 | -35.5344 | -32.3921
21.4137 | -35.535 | -32.3985
21.4199 | -35.5355 | -32.4056
21.426 | -35.536 | -32.4119
21.4321 | -35.5365 | -32.4184
21.4382 | -35.5371 | -32.4251
21.4443 | -35.5375 | -32.4317
21.4504 | -35.5381 | -32.4382
21.4565 | -35.5386 | -32.4445
21.4626 | -35.539 | -32.451
21.4687 | -35.5395 | -32.4575
21.4748 | -35.54 | -32.4642
21.4809 | -35.5405 | -32.4704
21.487 | -35.541 | -32.4771
21.4931 | -35.5415 | -32.484
21.4992 | -35.5421 | -32.4907
21.5053 | -35.5426 | -32.4975
21.5114 | -35.5431 | -32.5041
21.5175 | -35.5436 | -32.5105
21.5236 | -35.5441 | -32.517
21.5297 | -35.5447 | -32.5233
21.5358 | -35.5452 | -32.5298
21.5419 | -35.5457 | -32.5364
21.548 | -35.5462 | -32.543
21.5541 | -35.5467 | -32.5494
21.5602 | -35.5472 | -32.556
21.5664 | -35.5478 | -32.5624
21.5725 | -35.5483 | -32.5689
21.5786 | -35.5488 | -32.5755
21.5847 | -35.5493 | -32.5821
21.5908 | -35.5498 | -32.5885
21.5969 | -35.5504 | -32.5948
21.603 | -35.5509 | -32.6015
21.6091 | -35.5514 | -32.6081
21.6152 | -35.5519 | -32.6147
21.6213 | -35.5524 | -32.6215
21.6274 | -35.5529 | -32.628
21.6335 | -35.5535 | -32.6341
21.6396 | -35.554 | -32.6409
21.6457 | -35.5545 | -32.6473
21.6518 | -35.5549 | -32.6535
21.6579 | -35.5554 | -32.66
21.664 | -35.556 | -32.6665
21.6701 | -35.5565 | -32.6727
21.6762 | -35.5569 | -32.6791
21.6823 | -35.5574 | -32.6855
21.6884 | -35.558 | -32.6922
21.6945 | -35.5585 | -32.699
21.7006 | -35.5589 | -32.7054
21.7068 | -35.5594 | -32.7117
21.7129 | -35.56 | -32.7183
21.719 | -35.5605 | -32.7248
21.7251 | -35.561 | -32.7312
21.7312 | -35.5615 | -32.7377
21.7373 | -35.562 | -32.7438
21.7434 | -35.5625 | -32.7504
21.7495 | -35.5631 | -32.757
21.7556 | -35.5635 | -32.7635
21.7617 | -35.5641 | -32.7704
21.7678 | -35.5646 | -32.7767
21.7739 | -35.5651 | -32.783
21.78 | -35.5656 | -32.7895
21.7861 | -35.5661 | -32.7962
21.7922 | -35.5666 | -32.8027
21.7983 | -35.5671 | -32.8094
21.8044 | -35.5677 | -32.8159
21.8105 | -35.5681 | -32.8221
21.8166 | -35.5687 | -32.8287
21.8227 | -35.5692 | -32.8352
21.8288 | -35.5697 | -32.8417
21.8349 | -35.5702 | -32.8483
21.841 | -35.5707 | -32.8546
21.8471 | -35.5713 | -32.8608
21.8533 | -35.5718 | -32.8672
21.8594 | -35.5723 | -32.874
21.8655 | -35.5728 | -32.8803
21.8716 | -35.5733 | -32.8867
21.8777 | -35.5738 | -32.8934
21.8838 | -35.5743 | -32.8996
21.8899 | -35.5748 | -32.9059
21.896 | -35.5753 | -32.9126
21.9021 | -35.5758 | -32.9194
21.9082 | -35.5764 | -32.926
21.9143 | -35.5769 | -32.9321
21.9204 | -35.5774 | -32.9383
21.9265 | -35.5779 | -32.9449
21.9326 | -35.5784 | -32.9514
21.9387 | -35.579 | -32.9581
21.9448 | -35.5795 | -32.9644
21.9509 | -35.58 | -32.9709
21.957 | -35.5805 | -32.9775
21.9631 | -35.581 | -32.9838
21.9692 | -35.5815 | -32.9902
21.9753 | -35.582 | -32.9967
21.9814 | -35.5826 | -33.0032
21.9875 | -35.5832 | -33.0093
21.9937 | -35.5837 | -33.0158
21.9998 | -35.5842 | -33.0221
22.0059 | -35.5847 | -33.0284
22.012 | -35.5852 | -33.0347
22.0181 | -35.5857 | -33.0411
22.0242 | -35.5862 | -33.0479
22.0303 | -35.5867 | -33.0546
22.0364 | -35.5872 | -33.0612
22.0425 | -35.5877 | -33.0675
22.0486 | -35.5882 | -33.0739
22.0547 | -35.5888 | -33.0801
22.0608 | -35.5893 | -33.0867
22.0669 | -35.5899 | -33.093
22.073 | -35.5904 | -33.0994
22.0791 | -35.591 | -33.1057
22.0852 | -35.5915 | -33.1123
22.0913 | -35.592 | -33.1185
22.0974 | -35.5925 | -33.1248
22.1035 | -35.593 | -33.131
22.1096 | -35.5935 | -33.1376
22.1157 | -35.5941 | -33.1443
22.1218 | -35.5946 | -33.1507
22.1279 | -35.5951 | -33.1569
22.134 | -35.5956 | -33.1634
22.1402 | -35.5962 | -33.1698
22.1463 | -35.5967 | -33.1764
22.1524 | -35.5972 | -33.183
22.1585 | -35.5977 | -33.1892
22.1646 | -35.5983 | -33.1952
22.1707 | -35.5988 | -33.2016
22.1768 | -35.5993 | -33.208
22.1829 | -35.5997 | -33.2148
22.189 | -35.6002 | -33.2213
22.1951 | -35.6007 | -33.228
22.2012 | -35.6012 | -33.2344
22.2073 | -35.6017 | -33.2409
22.2134 | -35.6023 | -33.2469
22.2195 | -35.6028 | -33.2532
22.2256 | -35.6033 | -33.2592
22.2317 | -35.6039 | -33.2655
22.2378 | -35.6044 | -33.2718
22.2439 | -35.6049 | -33.2782
22.25 | -35.6054 | -33.2844
22.2561 | -35.6059 | -33.2907
22.2622 | -35.6065 | -33.2969
22.2683 | -35.607 | -33.3028
22.2744 | -35.6076 | -33.3091
22.2806 | -35.6081 | -33.3157
22.2867 | -35.6086 | -33.3219
22.2928 | -35.6092 | -33.3283
22.2989 | -35.6097 | -33.3346
22.305 | -35.6102 | -33.3409
22.3111 | -35.6108 | -33.3471
22.3172 | -35.6113 | -33.3537
22.3233 | -35.6118 | -33.3602
22.3294 | -35.6124 | -33.3667
22.3355 | -35.6129 | -33.3727
22.3416 | -35.6134 | -33.3788
22.3477 | -35.6139 | -33.3854
22.3538 | -35.6144 | -33.3916
22.3599 | -35.6149 | -33.3982
22.366 | -35.6155 | -33.4044
22.3721 | -35.616 | -33.4103
22.3782 | -35.6165 | -33.4166
22.3843 | -35.6171 | -33.4229
22.3904 | -35.6176 | -33.4294
22.3965 | -35.6181 | -33.436
22.4026 | -35.6187 | -33.4421
22.4087 | -35.6192 | -33.4482
22.4148 | -35.6197 | -33.4545
22.4209 | -35.6202 | -33.4608
22.4271 | -35.6207 | -33.4672
22.4332 | -35.6212 | -33.4735
22.4393 | -35.6218 | -33.4799
22.4454 | -35.6223 | -33.4861
22.4515 | -35.6228 | -33.4922
22.4576 | -35.6233 | -33.4983
22.4637 | -35.6239 | -33.5047
22.4698 | -35.6244 | -33.511
22.4759 | -35.6249 | -33.5171
22.482 | -35.6254 | -33.5236
22.4881 | -35.6259 | -33.53
22.4942 | -35.6264 | -33.5362
22.5003 | -35.6269 | -33.5423
22.5064 | -35.6275 | -33.5486
22.5125 | -35.628 | -33.5546
22.5186 | -35.6285 | -33.5609
22.5247 | -35.629 | -33.5669
22.5308 | -35.6296 | -33.5733
22.5369 | -35.6301 | -33.5796
22.543 | -35.6306 | -33.5859
22.5491 | -35.6311 | -33.5918
22.5552 | -35.6317 | -33.5983
22.5613 | -35.6322 | -33.6046
22.5675 | -35.6327 | -33.6109
22.5736 | -35.6332 | -33.6172
22.5797 | -35.6337 | -33.6236
22.5858 | -35.6343 | -33.6294
22.5919 | -35.6348 | -33.6355
22.598 | -35.6353 | -33.6418
22.6041 | -35.6358 | -33.648
22.6102 | -35.6363 | -33.6542
22.6163 | -35.6368 | -33.6603
22.6224 | -35.6374 | -33.6667
22.6285 | -35.6379 | -33.6731
22.6346 | -35.6384 | -33.6794
22.6407 | -35.6389 | -33.6856
22.6468 | -35.6395 | -33.692
22.6529 | -35.64 | -33.6981
22.659 | -35.6405 | -33.7045
22.6651 | -35.641 | -33.7107
22.6712 | -35.6415 | -33.7168
22.6773 | -35.6421 | -33.723
22.6834 | -35.6426 | -33.7293
22.6895 | -35.6432 | -33.7351
22.6956 | -35.6437 | -33.7414
22.7017 | -35.6442 | -33.7477
22.7079 | -35.6448 | -33.7541
22.714 | -35.6453 | -33.7604
22.7201 | -35.6458 | -33.7665
22.7262 | -35.6464 | -33.7722
22.7323 | -35.647 | -33.7785
22.7384 | -35.6475 | -33.7848
22.7445 | -35.648 | -33.7909
22.7506 | -35.6486 | -33.7971
22.7567 | -35.6491 | -33.8037
22.7628 | -35.6497 | -33.8095
22.7689 | -35.6502 | -33.8156
22.775 | -35.6507 | -33.8218
22.7811 | -35.6512 | -33.8281
22.7872 | -35.6517 | -33.8345
22.7933 | -35.6522 | -33.8405
22.7994 | -35.6528 | -33.8468
22.8055 | -35.6533 | -33.8531
22.8116 | -35.6538 | -33.8593
22.8177 | -35.6543 | -33.8657
22.8238 | -35.6549 | -33.8716
22.8299 | -35.6554 | -33.8775
22.836 | -35.6559 | -33.8838
22.8421 | -35.6565 | -33.8902
22.8482 | -35.657 | -33.8963
22.8544 | -35.6575 | -33.9024
22.8605 | -35.658 | -33.9087
22.8666 | -35.6585 | -33.9148
22.8727 | -35.659 | -33.9211
22.8788 | -35.6595 | -33.927
22.8849 | -35.66 | -33.9332
22.891 | -35.6605 | -33.9393
22.8971 | -35.6611 | -33.9456
22.9032 | -35.6616 | -33.9516
22.9093 | -35.6621 | -33.9578
22.9154 | -35.6627 | -33.9639
22.9215 | -35.6632 | -33.9701
22.9276 | -35.6637 | -33.9762
22.9337 | -35.6642 | -33.9818
22.9398 | -35.6647 | -33.9879
22.9459 | -35.6653 | -33.9943
22.952 | -35.6658 | -34.0005
22.9581 | -35.6663 | -34.0068
22.9642 | -35.6668 | -34.0132
22.9703 | -35.6674 | -34.019
22.9764 | -35.6679 | -34.0252
22.9825 | -35.6684 | -34.0313
22.9886 | -35.6689 | -34.0375
22.9948 | -35.6694 | -34.0435
23.0009 | -35.67 | -34.0497
23.007 | -35.6705 | -34.056
23.0131 | -35.6711 | -34.0623
23.0192 | -35.6716 | -34.0685
23.0253 | -35.6721 | -34.075
23.0314 | -35.6726 | -34.0813
23.0375 | -35.6731 | -34.0876
23.0436 | -35.6737 | -34.0935
23.0497 | -35.6742 | -34.0998
23.0558 | -35.6747 | -34.1059
23.0619 | -35.6753 | -34.1121
23.068 | -35.6758 | -34.1183
23.0741 | -35.6763 | -34.124
23.0802 | -35.6769 | -34.1303
23.0863 | -35.6774 | -34.1363
23.0924 | -35.6779 | -34.1425
23.0985 | -35.6784 | -34.1485
23.1046 | -35.6789 | -34.1547
23.1107 | -35.6795 | -34.1608
23.1168 | -35.68 | -34.167
23.1229 | -35.6805 | -34.1731
23.129 | -35.6811 | -34.1793
23.1351 | -35.6816 | -34.1854
23.1413 | -35.6821 | -34.1916
23.1474 | -35.6826 | -34.1978
23.1535 | -35.6832 | -34.2038
23.1596 | -35.6837 | -34.21
23.1657 | -35.6843 | -34.2164
23.1718 | -35.6848 | -34.2224
23.1779 | -35.6854 | -34.2287
23.184 | -35.686 | -34.2347
23.1901 | -35.6865 | -34.2409
23.1962 | -35.687 | -34.2472
23.2023 | -35.6875 | -34.2534
23.2084 | -35.6881 | -34.2597
23.2145 | -35.6886 | -34.2652
23.2206 | -35.6891 | -34.2716
23.2267 | -35.6896 | -34.2776
23.2328 | -35.6902 | -34.2836
23.2389 | -35.6907 | -34.2897
23.245 | -35.6912 | -34.2957
23.2511 | -35.6917 | -34.3015
23.2572 | -35.6923 | -34.3078
23.2633 | -35.6928 | -34.3139
23.2694 | -35.6933 | -34.3203
23.2755 | -35.6938 | -34.3265
23.2817 | -35.6943 | -34.3328
23.2878 | -35.6949 | -34.3385
23.2939 | -35.6954 | -34.3446
23.3 | -35.6959 | -34.3511
23.3061 | -35.6965 | -34.3574
23.3122 | -35.697 | -34.3634
23.3183 | -35.6975 | -34.3692
23.3244 | -35.6981 | -34.3748
23.3305 | -35.6986 | -34.381
23.3366 | -35.6991 | -34.387
23.3427 | -35.6996 | -34.3931
23.3488 | -35.7001 | -34.3994
23.3549 | -35.7006 | -34.4054
23.361 | -35.7011 | -34.4113
23.3671 | -35.7016 | -34.4172
23.3732 | -35.7021 | -34.4233
23.3793 | -35.7027 | -34.4294
23.3854 | -35.7032 | -34.4358
23.3915 | -35.7038 | -34.4419
23.3976 | -35.7043 | -34.448
23.4037 | -35.7048 | -34.4539
23.4098 | -35.7053 | -34.4601
23.4159 | -35.7059 | -34.4662
23.422 | -35.7064 | -34.4724
23.4282 | -35.707 | -34.4784
23.4343 | -35.7075 | -34.4843
23.4404 | -35.7081 | -34.4902
23.4465 | -35.7086 | -34.4962
23.4526 | -35.7092 | -34.5023
23.4587 | -35.7097 | -34.5081
23.4648 | -35.7102 | -34.514
23.4709 | -35.7108 | -34.5201
23.477 | -35.7113 | -34.526
23.4831 | -35.7118 | -34.5318
23.4892 | -35.7123 | -34.5376
23.4953 | -35.7128 | -34.5437
23.5014 | -35.7133 | -34.5495
23.5075 | -35.7138 | -34.5554
23.5136 | -35.7144 | -34.5616
23.5197 | -35.7149 | -34.5673
23.5258 | -35.7154 | -34.5734
23.5319 | -35.7159 | -34.5792
23.538 | -35.7164 | -34.585
23.5441 | -35.7169 | -34.591
23.5502 | -35.7174 | -34.5973
23.5563 | -35.718 | -34.6036
23.5624 | -35.7184 | -34.6098
23.5686 | -35.7189 | -34.6153
23.5747 | -35.7195 | -34.6211
23.5808 | -35.72 | -34.6269
23.5869 | -35.7205 | -34.6331
23.593 | -35.7211 | -34.639
23.5991 | -35.7216 | -34.6451
23.6052 | -35.7221 | -34.6508
23.6113 | -35.7227 | -34.6569
23.6174 | -35.7232 | -34.6628
23.6235 | -35.7237 | -34.6688
23.6296 | -35.7242 | -34.675
23.6357 | -35.7247 | -34.6811
23.6418 | -35.7252 | -34.6869
23.6479 | -35.7257 | -34.6928
23.654 | -35.7262 | -34.6988
23.6601 | -35.7268 | -34.7049
23.6662 | -35.7273 | -34.7108
23.6723 | -35.7278 | -34.7167
23.6784 | -35.7283 | -34.7227
23.6845 | -35.7288 | -34.7285
23.6906 | -35.7294 | -34.7347
23.6967 | -35.7299 | -34.7406
23.7028 | -35.7304 | -34.7465
23.7089 | -35.731 | -34.752
23.7151 | -35.7315 | -34.7582
23.7212 | -35.732 | -34.7639
23.7273 | -35.7325 | -34.7701
23.7334 | -35.7331 | -34.7762
23.7395 | -35.7336 | -34.7823
23.7456 | -35.734 | -34.7885
23.7517 | -35.7345 | -34.7943
23.7578 | -35.735 | -34.8005
23.7639 | -35.7355 | -34.8066
23.77 | -35.736 | -34.8127
23.7761 | -35.7365 | -34.8186
23.7822 | -35.737 | -34.8241
23.7883 | -35.7376 | -34.8298
23.7944 | -35.7381 | -34.8356
23.8005 | -35.7385 | -34.8417
23.8066 | -35.7391 | -34.8476
23.8127 | -35.7396 | -34.8534
23.8188 | -35.7402 | -34.8595
23.8249 | -35.7407 | -34.8655
23.831 | -35.7412 | -34.8716
23.8371 | -35.7418 | -34.8774
23.8432 | -35.7423 | -34.8833
23.8493 | -35.7428 | -34.8887
23.8555 | -35.7433 | -34.8947
23.8616 | -35.7438 | -34.9005
23.8677 | -35.7443 | -34.9068
23.8738 | -35.7448 | -34.9126
23.8799 | -35.7453 | -34.9186
23.886 | -35.7459 | -34.9242
23.8921 | -35.7464 | -34.9302
23.8982 | -35.7469 | -34.936
23.9043 | -35.7474 | -34.9419
23.9104 | -35.7479 | -34.9477
23.9165 | -35.7484 | -34.9537
23.9226 | -35.7489 | -34.9595
23.9287 | -35.7494 | -34.9656
23.9348 | -35.7499 | -34.9717
23.9409 | -35.7504 | -34.9777
23.947 | -35.7509 | -34.9838
23.9531 | -35.7514 | -34.9893
23.9592 | -35.752 | -34.9951
23.9653 | -35.7525 | -35.0009
23.9714 | -35.753 | -35.007
23.9775 | -35.7535 | -35.0129
23.9836 | -35.754 | -35.0189
23.9897 | -35.7545 | -35.0245
23.9958 | -35.7551 | -35.0304
24.002 | -35.7555 | -35.0365
24.0081 | -35.7561 | -35.0423
24.0142 | -35.7566 | -35.0482
24.0203 | -35.7571 | -35.0543
24.0264 | -35.7576 | -35.0599
24.0325 | -35.7581 | -35.0657
24.0386 | -35.7586 | -35.0717
24.0447 | -35.7591 | -35.0777
24.0508 | -35.7596 | -35.0836
24.0569 | -35.7601 | -35.0897
24.063 | -35.7607 | -35.0951
24.0691 | -35.7612 | -35.1009
24.0752 | -35.7617 | -35.1069
24.0813 | -35.7622 | -35.1126
24.0874 | -35.7628 | -35.1182
24.0935 | -35.7633 | -35.1241
24.0996 | -35.7638 | -35.1297
24.1057 | -35.7643 | -35.1356
24.1118 | -35.7647 | -35.1417
24.1179 | -35.7653 | -35.1478
24.124 | -35.7658 | -35.1536
24.1301 | -35.7663 | -35.159
24.1362 | -35.7668 | -35.165
24.1424 | -35.7673 | -35.1708
24.1485 | -35.7678 | -35.1768
24.1546 | -35.7684 | -35.1828
24.1607 | -35.7688 | -35.1888
24.1668 | -35.7693 | -35.1944
24.1729 | -35.7698 | -35.2002
24.179 | -35.7704 | -35.2061
24.1851 | -35.7708 | -35.2116
24.1912 | -35.7714 | -35.2175
24.1973 | -35.7719 | -35.2236
24.2034 | -35.7724 | -35.2294
24.2095 | -35.7729 | -35.2354
24.2156 | -35.7734 | -35.2412
24.2217 | -35.774 | -35.2473
24.2278 | -35.7745 | -35.2532
24.2339 | -35.775 | -35.259
24.24 | -35.7755 | -35.2644
24.2461 | -35.776 | -35.2701
24.2522 | -35.7765 | -35.2761
24.2583 | -35.777 | -35.2821
24.2644 | -35.7775 | -35.288
24.2705 | -35.778 | -35.2939
24.2766 | -35.7786 | -35.2997
24.2827 | -35.779 | -35.3055
24.2889 | -35.7796 | -35.3114
24.295 | -35.7801 | -35.3173
24.3011 | -35.7806 | -35.3228
24.3072 | -35.7811 | -35.3285
24.3133 | -35.7816 | -35.3344
24.3194 | -35.7821 | -35.3403
24.3255 | -35.7826 | -35.3462
24.3316 | -35.7831 | -35.3517
24.3377 | -35.7836 | -35.3575
24.3438 | -35.7841 | -35.3631
24.3499 | -35.7846 | -35.3688
24.356 | -35.785 | -35.3747
24.3621 | -35.7855 | -35.3805
24.3682 | -35.786 | -35.3864
24.3743 | -35.7865 | -35.3925
24.3804 | -35.787 | -35.3981
24.3865 | -35.7875 | -35.4041
24.3926 | -35.788 | -35.4099
24.3987 | -35.7885 | -35.4156
24.4048 | -35.789 | -35.421
24.4109 | -35.7895 | -35.4266
24.417 | -35.79 | -35.4324
24.4231 | -35.7905 | -35.438
24.4293 | -35.791 | -35.444
24.4354 | -35.7915 | -35.4498
24.4415 | -35.792 | -35.4554
24.4476 | -35.7925 | -35.4613
24.4537 | -35.793 | -35.4672
24.4598 | -35.7935 | -35.4731
24.4659 | -35.7941 | -35.4789
24.472 | -35.7946 | -35.4848
24.4781 | -35.7951 | -35.4906
24.4842 | -35.7956 | -35.4961
24.4903 | -35.7961 | -35.5019
24.4964 | -35.7966 | -35.5075
24.5025 | -35.797 | -35.5133
24.5086 | -35.7976 | -35.5192
24.5147 | -35.7981 | -35.5249
24.5208 | -35.7986 | -35.5303
24.5269 | -35.7991 | -35.5361
24.533 | -35.7997 | -35.5417
24.5391 | -35.8001 | -35.5476
24.5452 | -35.8006 | -35.5534
24.5513 | -35.8011 | -35.5592
24.5574 | -35.8017 | -35.5653
24.5635 | -35.8021 | -35.571
24.5696 | -35.8026 | -35.5767
24.5758 | -35.8032 | -35.5823
24.5819 | -35.8037 | -35.5881
24.588 | -35.8042 | -35.5938
24.5941 | -35.8047 | -35.5996
24.6002 | -35.8052 | -35.6052
24.6063 | -35.8057 | -35.6111
24.6124 | -35.8062 | -35.6167
24.6185 | -35.8068 | -35.6225
24.6246 | -35.8072 | -35.6285
24.6307 | -35.8077 | -35.6344
24.6368 | -35.8082 | -35.6403
24.6429 | -35.8087 | -35.6461
24.649 | -35.8092 | -35.6519
24.6551 | -35.8097 | -35.6575
24.6612 | -35.8102 | -35.6632
24.6673 | -35.8107 | -35.6689
24.6734 | -35.8112 | -35.6746
24.6795 | -35.8116 | -35.6803
24.6856 | -35.8121 | -35.6862
24.6917 | -35.8127 | -35.6919
24.6978 | -35.8132 | -35.6973
24.7039 | -35.8137 | -35.7031
24.71 | -35.8142 | -35.7086
24.7162 | -35.8147 | -35.7144
24.7223 | -35.8152 | -35.72
24.7284 | -35.8157 | -35.7255
24.7345 | -35.8162 | -35.7313
24.7406 | -35.8167 | -35.737
24.7467 | -35.8172 | -35.7427
24.7528 | -35.8177 | -35.7485
24.7589 | -35.8182 | -35.7541
24.765 | -35.8187 | -35.7598
24.7711 | -35.8192 | -35.7656
24.7772 | -35.8197 | -35.7713
24.7833 | -35.8202 | -35.7772
24.7894 | -35.8206 | -35.7828
24.7955 | -35.8211 | -35.7885
24.8016 | -35.8217 | -35.7939
24.8077 | -35.8221 | -35.7997
24.8138 | -35.8226 | -35.8053
24.8199 | -35.8231 | -35.811
24.826 | -35.8236 | -35.8167
24.8321 | -35.8241 | -35.8225
24.8382 | -35.8246 | -35.828
24.8443 | -35.825 | -35.8336
24.8504 | -35.8256 | -35.8393
24.8565 | -35.8261 | -35.8453
24.8627 | -35.8265 | -35.851
24.8688 | -35.827 | -35.8564
24.8749 | -35.8275 | -35.8621
24.881 | -35.828 | -35.868
24.8871 | -35.8286 | -35.874
24.8932 | -35.8291 | -35.88
24.8993 | -35.8295 | -35.8857
24.9054 | -35.83 | -35.8913
24.9115 | -35.8305 | -35.897
24.9176 | -35.8311 | -35.9028
24.9237 | -35.8316 | -35.9087
24.9298 | -35.8321 | -35.9145
24.9359 | -35.8326 | -35.9205
24.942 | -35.833 | -35.926
24.9481 | -35.8335 | -35.9318
24.9542 | -35.834 | -35.9377
24.9603 | -35.8345 | -35.9434
24.9664 | -35.835 | -35.9489
24.9725 | -35.8355 | -35.9546
24.9786 | -35.8359 | -35.9599
24.9847 | -35.8364 | -35.9656
24.9908 | -35.8369 | -35.9716
24.9969 | -35.8374 | -35.9772
25.0031 | -35.8379 | -35.9831
25.0092 | -35.8384 | -35.9886
25.0153 | -35.8388 | -35.9941
25.0214 | -35.8393 | -35.9999
25.0275 | -35.8398 | -36.0056
25.0336 | -35.8403 | -36.0112
25.0397 | -35.8408 | -36.0169
25.0458 | -35.8413 | -36.0223
25.0519 | -35.8417 | -36.028
25.058 | -35.8422 | -36.0337
25.0641 | -35.8427 | -36.0395
25.0702 | -35.8432 | -36.0451
25.0763 | -35.8437 | -36.0508
25.0824 | -35.8442 | -36.0562
25.0885 | -35.8447 | -36.0617
25.0946 | -35.8452 | -36.0675
25.1007 | -35.8457 | -36.0731
25.1068 | -35.8462 | -36.0787
25.1129 | -35.8466 | -36.0843
25.119 | -35.8471 | -36.0899
25.1251 | -35.8476 | -36.0956
25.1312 | -35.8481 | -36.1012
25.1373 | -35.8486 | -36.107
25.1435 | -35.8492 | -36.1128
25.1496 | -35.8496 | -36.1184
25.1557 | -35.8501 | -36.124
25.1618 | -35.8506 | -36.1296
25.1679 | -35.8511 | -36.1354
25.174 | -35.8516 | -36.1408
25.1801 | -35.852 | -36.1465
25.1862 | -35.8525 | -36.152
25.1923 | -35.853 | -36.1575
25.1984 | -35.8535 | -36.1632
25.2045 | -35.854 | -36.169
25.2106 | -35.8545 | -36.1746
25.2167 | -35.855 | -36.1804
25.2228 | -35.8554 | -36.1859
25.2289 | -35.8559 | -36.1915
25.235 | -35.8564 | -36.1973
25.2411 | -35.8568 | -36.203
25.2472 | -35.8573 | -36.2088
25.2533 | -35.8578 | -36.2145
25.2594 | -35.8582 | -36.22
25.2655 | -35.8587 | -36.2255
25.2716 | -35.8592 | -36.2309
25.2777 | -35.8596 | -36.2363
25.2838 | -35.8601 | -36.2421
25.29 | -35.8605 | -36.2474
25.2961 | -35.861 | -36.2526
25.3022 | -35.8615 | -36.2581
25.3083 | -35.862 | -36.2636
25.3144 | -35.8625 | -36.2693
25.3205 | -35.8629 | -36.275
25.3266 | -35.8634 | -36.2804
25.3327 | -35.8639 | -36.2861
25.3388 | -35.8644 | -36.2916
25.3449 | -35.8649 | -36.2973
25.351 | -35.8654 | -36.303
25.3571 | -35.8659 | -36.3087
25.3632 | -35.8663 | -36.3141
25.3693 | -35.8668 | -36.3196
25.3754 | -35.8673 | -36.3252
25.3815 | -35.8677 | -36.3306
25.3876 | -35.8682 | -36.336
25.3937 | -35.8687 | -36.3418
25.3998 | -35.8692 | -36.3472
25.4059 | -35.8697 | -36.3529
25.412 | -35.8702 | -36.3583
25.4181 | -35.8706 | -36.3638
25.4242 | -35.8711 | -36.3695
25.4304 | -35.8716 | -36.3751
25.4365 | -35.872 | -36.3807
25.4426 | -35.8725 | -36.3864
25.4487 | -35.873 | -36.3922
25.4548 | -35.8734 | -36.398
25.4609 | -35.8739 | -36.4037
25.467 | -35.8744 | -36.4094
25.4731 | -35.8748 | -36.4151
25.4792 | -35.8753 | -36.4208
25.4853 | -35.8759 | -36.4264
25.4914 | -35.8764 | -36.4321
25.4975 | -35.8769 | -36.4378
25.5036 | -35.8773 | -36.4435
25.5097 | -35.8778 | -36.4489
25.5158 | -35.8783 | -36.4541
25.5219 | -35.8788 | -36.4597
25.528 | -35.8792 | -36.4653
25.5341 | -35.8798 | -36.471
25.5402 | -35.8802 | -36.4764
25.5463 | -35.8807 | -36.4821
25.5524 | -35.8812 | -36.4876
25.5585 | -35.8817 | -36.4932
25.5646 | -35.8822 | -36.4985
25.5707 | -35.8826 | -36.5041
25.5769 | -35.8832 | -36.5097
25.583 | -35.8837 | -36.5153
25.5891 | -35.8842 | -36.5207
25.5952 | -35.8846 | -36.5262
25.6013 | -35.8851 | -36.5316
25.6074 | -35.8856 | -36.5371
25.6135 | -35.8861 | -36.5427
25.6196 | -35.8865 | -36.5485
25.6257 | -35.887 | -36.5543
25.6318 | -35.8875 | -36.5599
25.6379 | -35.8879 | -36.5657
25.644 | -35.8884 | -36.5715
25.6501 | -35.8889 | -36.577
25.6562 | -35.8894 | -36.5824
25.6623 | -35.8899 | -36.5878
25.6684 | -35.8903 | -36.5933
25.6745 | -35.8909 | -36.5987
25.6806 | -35.8913 | -36.6041
25.6867 | -35.8918 | -36.6098
25.6928 | -35.8923 | -36.6154
25.6989 | -35.8927 | -36.621
25.705 | -35.8932 | -36.6265
25.7111 | -35.8937 | -36.632
25.7173 | -35.8942 | -36.6375
25.7234 | -35.8947 | -36.6429
25.7295 | -35.8952 | -36.6482
25.7356 | -35.8956 | -36.654
25.7417 | -35.8961 | -36.6595
25.7478 | -35.8965 | -36.6648
25.7539 | -35.897 | -36.6701
25.76 | -35.8975 | -36.6757
25.7661 | -35.898 | -36.6813
25.7722 | -35.8984 | -36.6869
25.7783 | -35.8989 | -36.6922
25.7844 | -35.8993 | -36.6974
25.7905 | -35.8998 | -36.7029
25.7966 | -35.9003 | -36.7085
25.8027 | -35.9007 | -36.7141
25.8088 | -35.9012 | -36.7195
25.8149 | -35.9016 | -36.7248
25.821 | -35.9021 | -36.7302
25.8271 | -35.9026 | -36.7357
25.8332 | -35.903 | -36.7412
25.8393 | -35.9035 | -36.7466
25.8454 | -35.904 | -36.7523
25.8515 | -35.9044 | -36.758
25.8576 | -35.9049 | -36.7636
25.8638 | -35.9054 | -36.7692
25.8699 | -35.9059 | -36.7749
25.876 | -35.9064 | -36.7804
25.8821 | -35.9069 | -36.7858
25.8882 | -35.9073 | -36.7911
25.8943 | -35.9078 | -36.7968
25.9004 | -35.9083 | -36.8024
25.9065 | -35.9088 | -36.8083
25.9126 | -35.9092 | -36.8139
25.9187 | -35.9097 | -36.8195
25.9248 | -35.9101 | -36.825
25.9309 | -35.9106 | -36.8305
25.937 | -35.9111 | -36.836
25.9431 | -35.9116 | -36.8415
25.9492 | -35.9121 | -36.8469
25.9553 | -35.9126 | -36.8524
25.9614 | -35.9131 | -36.8578
25.9675 | -35.9135 | -36.8634
25.9736 | -35.914 | -36.869
25.9797 | -35.9145 | -36.8746
25.9858 | -35.915 | -36.88
25.9919 | -35.9154 | -36.8857
25.998 | -35.9159 | -36.8911
26.0042 | -35.9163 | -36.8968
26.0103 | -35.9168 | -36.9023
26.0164 | -35.9172 | -36.9077
26.0225 | -35.9177 | -36.9133
26.0286 | -35.9182 | -36.9189
26.0347 | -35.9187 | -36.9246
26.0408 | -35.9191 | -36.9302
26.0469 | -35.9195 | -36.9355
26.053 | -35.92 | -36.9411
26.0591 | -35.9205 | -36.9465
26.0652 | -35.9209 | -36.9519
26.0713 | -35.9214 | -36.9573
26.0774 | -35.9218 | -36.9629
26.0835 | -35.9223 | -36.9685
26.0896 | -35.9228 | -36.9743
26.0957 | -35.9233 | -36.9798
26.1018 | -35.9237 | -36.9853
26.1079 | -35.9242 | -36.9909
26.114 | -35.9246 | -36.9963
26.1201 | -35.925 | -37.0019
26.1262 | -35.9255 | -37.0073
26.1323 | -35.926 | -37.013
26.1384 | -35.9264 | -37.0178
26.1445 | -35.9269 | -37.0232
26.1507 | -35.9274 | -37.0286
26.1568 | -35.9279 | -37.0342
26.1629 | -35.9283 | -37.0397
26.169 | -35.9287 | -37.0449
26.1751 | -35.9292 | -37.0504
26.1812 | -35.9297 | -37.0558
26.1873 | -35.9301 | -37.0613
26.1934 | -35.9306 | -37.0669
26.1995 | -35.931 | -37.0724
26.2056 | -35.9314 | -37.0778
26.2117 | -35.9318 | -37.0832
26.2178 | -35.9323 | -37.0882
26.2239 | -35.9327 | -37.0937
26.23 | -35.9332 | -37.0993
26.2361 | -35.9336 | -37.1047
26.2422 | -35.934 | -37.1099
26.2483 | -35.9345 | -37.1152
26.2544 | -35.9349 | -37.1205
26.2605 | -35.9354 | -37.1261
26.2666 | -35.9359 | -37.1318
26.2727 | -35.9363 | -37.1373
26.2788 | -35.9368 | -37.1427
26.2849 | -35.9373 | -37.1481
26.2911 | -35.9378 | -37.1535
26.2972 | -35.9382 | -37.159
26.3033 | -35.9387 | -37.1644
26.3094 | -35.9392 | -37.1696
26.3155 | -35.9396 | -37.1749
26.3216 | -35.94 | -37.1807
26.3277 | -35.9405 | -37.1861
26.3338 | -35.941 | -37.1916
26.3399 | -35.9415 | -37.1971
26.346 | -35.942 | -37.2021
26.3521 | -35.9424 | -37.2075
26.3582 | -35.9429 | -37.2129
26.3643 | -35.9433 | -37.2185
26.3704 | -35.9437 | -37.224
26.3765 | -35.9442 | -37.2293
26.3826 | -35.9446 | -37.2344
26.3887 | -35.9451 | -37.2398
26.3948 | -35.9455 | -37.2454
26.4009 | -35.946 | -37.2508
26.407 | -35.9464 | -37.2565
26.4131 | -35.9469 | -37.262
26.4192 | -35.9473 | -37.2673
26.4253 | -35.9478 | -37.2727
26.4314 | -35.9483 | -37.2784
26.4376 | -35.9487 | -37.2838
26.4437 | -35.9491 | -37.2891
26.4498 | -35.9496 | -37.2948
26.4559 | -35.9501 | -37.2996
26.462 | -35.9505 | -37.3049
26.4681 | -35.951 | -37.3103
26.4742 | -35.9515 | -37.3156
26.4803 | -35.9519 | -37.3208
26.4864 | -35.9524 | -37.3261
26.4925 | -35.9528 | -37.3314
26.4986 | -35.9532 | -37.3366
26.5047 | -35.9537 | -37.3418
26.5108 | -35.9541 | -37.3474
26.5169 | -35.9546 | -37.3529
26.523 | -35.955 | -37.3579
26.5291 | -35.9554 | -37.3633
26.5352 | -35.9559 | -37.3684
26.5413 | -35.9563 | -37.3739
26.5474 | -35.9568 | -37.3797
26.5535 | -35.9572 | -37.3853
26.5596 | -35.9576 | -37.3904
26.5657 | -35.958 | -37.3958
26.5718 | -35.9585 | -37.4013
26.578 | -35.959 | -37.4066
26.5841 | -35.9594 | -37.412
26.5902 | -35.9598 | -37.4172
26.5963 | -35.9602 | -37.4223
26.6024 | -35.9606 | -37.4281
26.6085 | -35.9611 | -37.4337
26.6146 | -35.9616 | -37.4392
26.6207 | -35.962 | -37.4451
26.6268 | -35.9625 | -37.4507
26.6329 | -35.9629 | -37.4558
26.639 | -35.9633 | -37.461
26.6451 | -35.9638 | -37.4664
26.6512 | -35.9642 | -37.4718
26.6573 | -35.9647 | -37.4774
26.6634 | -35.9651 | -37.4826
26.6695 | -35.9655 | -37.4885
26.6756 | -35.966 | -37.4942
26.6817 | -35.9664 | -37.4998
26.6878 | -35.9669 | -37.505
26.6939 | -35.9674 | -37.5104
26.7 | -35.9679 | -37.5158
26.7061 | -35.9683 | -37.5203
26.7122 | -35.9687 | -37.5258
26.7183 | -35.9692 | -37.5309
26.7245 | -35.9697 | -37.536
26.7306 | -35.97 | -37.5419
26.7367 | -35.9704 | -37.5471
26.7428 | -35.9709 | -37.5523
26.7489 | -35.9714 | -37.5579
26.755 | -35.9718 | -37.5634
26.7611 | -35.9723 | -37.5692
26.7672 | -35.9727 | -37.5743
26.7733 | -35.9732 | -37.5795
26.7794 | -35.9736 | -37.5851
26.7855 | -35.9741 | -37.59
26.7916 | -35.9745 | -37.5957
26.7977 | -35.975 | -37.601
26.8038 | -35.9755 | -37.6062
26.8099 | -35.976 | -37.6119
26.816 | -35.9764 | -37.6175
26.8221 | -35.9769 | -37.623
26.8282 | -35.9774 | -37.6283
26.8343 | -35.9778 | -37.6337
26.8404 | -35.9782 | -37.6393
26.8465 | -35.9787 | -37.6446
26.8526 | -35.9791 | -37.6501
26.8587 | -35.9796 | -37.6556
26.8649 | -35.9801 | -37.6614
26.871 | -35.9805 | -37.6669
26.8771 | -35.9809 | -37.6722
26.8832 | -35.9814 | -37.6778
26.8893 | -35.9818 | -37.6828
26.8954 | -35.9822 | -37.6882
26.9015 | -35.9827 | -37.6933
26.9076 | -35.9831 | -37.6989
26.9137 | -35.9835 | -37.704
26.9198 | -35.9839 | -37.7096
26.9259 | -35.9843 | -37.7152
26.932 | -35.9848 | -37.7204
26.9381 | -35.9853 | -37.7258
26.9442 | -35.9857 | -37.7307
26.9503 | -35.9861 | -37.7358
26.9564 | -35.9866 | -37.741
26.9625 | -35.9871 | -37.7465
26.9686 | -35.9875 | -37.7523
26.9747 | -35.9881 | -37.7574
26.9808 | -35.9885 | -37.7625
26.9869 | -35.989 | -37.7678
26.993 | -35.9895 | -37.7733
26.9991 | -35.9899 | -37.7781
27.0052 | -35.9903 | -37.7831
27.0114 | -35.9907 | -37.7885
27.0175 | -35.9911 | -37.7938
27.0236 | -35.9915 | -37.7991
27.0297 | -35.9919 | -37.8044
27.0358 | -35.9923 | -37.8096
27.0419 | -35.9927 | -37.8148
27.048 | -35.9932 | -37.8205
27.0541 | -35.9937 | -37.826
27.0602 | -35.9941 | -37.8312
27.0663 | -35.9945 | -37.8366
27.0724 | -35.995 | -37.8417
27.0785 | -35.9954 | -37.8469
27.0846 | -35.9959 | -37.8522
27.0907 | -35.9963 | -37.8573
27.0968 | -35.9968 | -37.8624
27.1029 | -35.9972 | -37.8679
27.109 | -35.9976 | -37.8733
27.1151 | -35.9981 | -37.8786
27.1212 | -35.9984 | -37.884
27.1273 | -35.9989 | -37.8896
27.1334 | -35.9994 | -37.8951
27.1395 | -35.9999 | -37.9009
27.1456 | -36.0003 | -37.9061
27.1518 | -36.0007 | -37.9115
27.1579 | -36.0012 | -37.9166
27.164 | -36.0016 | -37.9218
27.1701 | -36.0021 | -37.9272
27.1762 | -36.0025 | -37.9325
27.1823 | -36.0028 | -37.9381
27.1884 | -36.0032 | -37.9434
27.1945 | -36.0038 | -37.9488
27.2006 | -36.0043 | -37.9545
27.2067 | -36.0047 | -37.9598
27.2128 | -36.0052 | -37.965
27.2189 | -36.0056 | -37.9701
27.225 | -36.006 | -37.9757
27.2311 | -36.0064 | -37.9808
27.2372 | -36.0069 | -37.9861
27.2433 | -36.0073 | -37.9915
27.2494 | -36.0078 | -37.997
27.2555 | -36.0083 | -38.0024
27.2616 | -36.0087 | -38.0079
27.2677 | -36.0091 | -38.013
27.2738 | -36.0095 | -38.0185
27.2799 | -36.0099 | -38.0239
27.286 | -36.0103 | -38.0289
27.2921 | -36.0107 | -38.0344
27.2983 | -36.011 | -38.0395
27.3044 | -36.0115 | -38.0448
27.3105 | -36.0119 | -38.0502
27.3166 | -36.0123 | -38.0551
27.3227 | -36.0127 | -38.0606
27.3288 | -36.0132 | -38.0659
27.3349 | -36.0137 | -38.0709
27.341 | -36.0141 | -38.076
27.3471 | -36.0146 | -38.0811
27.3532 | -36.015 | -38.0861
27.3593 | -36.0154 | -38.0914
27.3654 | -36.0159 | -38.0969
27.3715 | -36.0163 | -38.102
27.3776 | -36.0167 | -38.1072
27.3837 | -36.0172 | -38.1125
27.3898 | -36.0176 | -38.1176
27.3959 | -36.018 | -38.1229
27.402 | -36.0184 | -38.1281
27.4081 | -36.0189 | -38.1332
27.4142 | -36.0193 | -38.1384
27.4203 | -36.0196 | -38.1434
27.4264 | -36.0201 | -38.1488
27.4325 | -36.0205 | -38.1542
27.4387 | -36.021 | -38.1588
27.4448 | -36.0215 | -38.1635
27.4509 | -36.0219 | -38.1692
27.457 | -36.0223 | -38.1744
27.4631 | -36.0227 | -38.1798
27.4692 | -36.0232 | -38.1853
27.4753 | -36.0236 | -38.1901
27.4814 | -36.024 | -38.1953
27.4875 | -36.0244 | -38.2006
27.4936 | -36.0248 | -38.206
27.4997 | -36.0253 | -38.2108
27.5058 | -36.0256 | -38.2161
27.5119 | -36.0261 | -38.2218
27.518 | -36.0266 | -38.2266
27.5241 | -36.027 | -38.2321
27.5302 | -36.0275 | -38.2371
27.5363 | -36.0279 | -38.2426
27.5424 | -36.0282 | -38.2472
27.5485 | -36.0286 | -38.2527
27.5546 | -36.0291 | -38.2581
27.5607 | -36.0295 | -38.2635
27.5668 | -36.0299 | -38.2686
27.5729 | -36.0303 | -38.274
27.5791 | -36.0307 | -38.279
27.5852 | -36.0311 | -38.2845
27.5913 | -36.0315 | -38.2899
27.5974 | -36.032 | -38.2953
27.6035 | -36.0324 | -38.3008
27.6096 | -36.0328 | -38.3056
27.6157 | -36.0332 | -38.3095
27.6218 | -36.0335 | -38.3145
27.6279 | -36.0339 | -38.3189
27.634 | -36.0343 | -38.3246
27.6401 | -36.0348 | -38.3299
27.6462 | -36.0353 | -38.3355
27.6523 | -36.0358 | -38.3412
27.6584 | -36.0362 | -38.3467
27.6645 | -36.0367 | -38.3517
27.6706 | -36.0371 | -38.3572
27.6767 | -36.0375 | -38.3623
27.6828 | -36.038 | -38.3677
27.6889 | -36.0383 | -38.373
27.695 | -36.0388 | -38.3781
27.7011 | -36.0392 | -38.3825
27.7072 | -36.0395 | -38.3879
27.7133 | -36.0399 | -38.3937
27.7194 | -36.0404 | -38.3989
27.7256 | -36.0408 | -38.4042
27.7317 | -36.0411 | -38.4087
27.7378 | -36.0415 | -38.4139
27.7439 | -36.042 | -38.4191
27.75 | -36.0424 | -38.4249
27.7561 | -36.0428 | -38.4296
27.7622 | -36.0432 | -38.4347
27.7683 | -36.0436 | -38.4397
27.7744 | -36.044 | -38.445
27.7805 | -36.0444 | -38.4504
27.7866 | -36.0448 | -38.4552
27.7927 | -36.0452 | -38.4601
27.7988 | -36.0456 | -38.4657
27.8049 | -36.046 | -38.4708
27.811 | -36.0463 | -38.4764
27.8171 | -36.0467 | -38.4817
27.8232 | -36.0471 | -38.4869
27.8293 | -36.0475 | -38.4917
27.8354 | -36.0478 | -38.4969
27.8415 | -36.0483 | -38.5022
27.8476 | -36.0487 | -38.5078
27.8537 | -36.049 | -38.5131
27.8598 | -36.0495 | -38.5177
27.866 | -36.05 | -38.5226
27.8721 | -36.0504 | -38.5281
27.8782 | -36.0509 | -38.5331
27.8843 | -36.0512 | -38.5382
27.8904 | -36.0516 | -38.5432
27.8965 | -36.052 | -38.548
27.9026 | -36.0523 | -38.5529
27.9087 | -36.0528 | -38.5578
27.9148 | -36.0532 | -38.5627
27.9209 | -36.0536 | -38.5684
27.927 | -36.0541 | -38.5739
27.9331 | -36.0546 | -38.5788
27.9392 | -36.0549 | -38.5842
27.9453 | -36.0553 | -38.5894
27.9514 | -36.0558 | -38.5949
27.9575 | -36.0561 | -38.6002
27.9636 | -36.0565 | -38.6054
27.9697 | -36.0569 | -38.6108
27.9758 | -36.0573 | -38.6165
27.9819 | -36.0577 | -38.6221
27.988 | -36.058 | -38.6268
27.9941 | -36.0583 | -38.6321
28.0002 | -36.0586 | -38.6373
28.0063 | -36.0589 | -38.6428
28.0125 | -36.0594 | -38.648
28.0186 | -36.0598 | -38.6531
28.0247 | -36.0602 | -38.658
28.0308 | -36.0607 | -38.6634
28.0369 | -36.0611 | -38.6682
28.043 | -36.0614 | -38.6736
28.0491 | -36.0618 | -38.6788
28.0552 | -36.0623 | -38.6838
28.0613 | -36.0626 | -38.6891
28.0674 | -36.063 | -38.6942
28.0735 | -36.0634 | -38.6997

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.7149 | -43.7602
21.4908 | -35.6528 | -35.8255
16.4948 | -34.9862 | -28.2977
12.6602 | -34.6 | -21.8931
9.7171 | -34.3661 | -16.8866
7.4582 | -34.2063 | -12.8299
5.7244 | -34.0955 | -9.5388
4.3936 | -34.0199 | -6.8902
3.3722 | -34.0076 | -4.799
2.5883 | -34.0017 | -3.2111
1.9866 | -34.0054 | -1.9456
1.5248 | -34.0039 | -0.931
1.1703 | -34.0165 | -0.2668
1 | -34.0159 | 0.063271
0.89822 | -34.0189 | 0.33664
0.68941 | -34.0034 | 0.92655
0.52919 | -33.9906 | 1.3863
0.40612 | -34.0047 | 1.9935
0.31171 | -33.9913 | 2.4496
0.23928 | -33.9963 | 3.1782
0.18358 | -34.0097 | 3.5882
0.14091 | -34.0283 | 4.4593
0.10819 | -34.0424 | 5.6205
0.083017 | -34.0644 | 7.0959
0.063705 | -34.1037 | 8.9715
0.048923 | -34.1768 | 11.4561
0.037479 | -34.2721 | 14.6343
0.028801 | -34.451 | 18.6959
0.022078 | -34.7718 | 24.1325
0.017023 | -35.2578 | 30.5433
0.013018 | -36.0639 | 38.2046
0.010014 | -37.1807 | 46.135

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003719 | -3.9735
0.007438 | -29.1104
0.011158 | -36.6014
0.014877 | -40.4883
0.018596 | -43.4711
0.022315 | -45.9683
0.026034 | -46.7428
0.029754 | -48.851
0.033473 | -50.35
0.037192 | -50.8547
0.040911 | -52.0003
0.04463 | -52.6354
0.04835 | -53.778
0.052069 | -55.3869
0.055788 | -57.5857
0.059507 | -56.1462
0.063226 | -56.802
0.066946 | -57.8582
0.070665 | -59.1599
0.074384 | -59.4568
0.078103 | -60.4109
0.081822 | -60.8602
0.085542 | -60.3883
0.089261 | -60.6039
0.09298 | -61.9647
0.096699 | -63.1958
0.10042 | -63.0289
0.10414 | -62.9222
0.10786 | -63.776
0.11158 | -63.8511
0.11529 | -64.2569
0.11901 | -64.1805
0.12273 | -64.6269
0.12645 | -66.0041
0.13017 | -66.0038
0.13389 | -65.7646
0.13761 | -65.7284
0.14133 | -67.0224
0.14505 | -68.8793
0.14877 | -68.6764
0.15249 | -67.8558
0.15621 | -66.9959
0.15993 | -66.9567
0.16365 | -68.3713
0.16736 | -68.6892
0.17108 | -68.6564
0.1748 | -69.1567
0.17852 | -69.922
0.18224 | -69.1908
0.18596 | -69.3066
0.18968 | -70.4887
0.1934 | -70.3732
0.19712 | -70.9766
0.20084 | -71.0315
0.20456 | -71.2703
0.20827 | -70.9168
0.21199 | -71.5228
0.21571 | -71.0303
0.21943 | -70.9433
0.22315 | -71.4062
0.22687 | -71.6918
0.23059 | -71.8426
0.23431 | -73.3777
0.23803 | -72.482
0.24175 | -72.3725
0.24547 | -72.7897
0.24919 | -74.1557
0.25291 | -72.5801
0.25662 | -72.2991
0.26034 | -72.9449
0.26406 | -74.066
0.26778 | -73.8308
0.2715 | -74.3464
0.27522 | -73.4985
0.27894 | -73.095
0.28266 | -74.3424
0.28638 | -75.6368
0.2901 | -74.5235
0.29382 | -74.6106
0.29754 | -75.6409
0.30125 | -75.7554
0.30497 | -74.8511
0.30869 | -73.8734
0.31241 | -74.2107
0.31613 | -74.6673
0.31985 | -74.4408
0.32357 | -74.536
0.32729 | -75.7231
0.33101 | -76.4049
0.33473 | -75.7156
0.33845 | -74.9904
0.34217 | -75.7662
0.34589 | -75.2538
0.3496 | -76.4969
0.35332 | -77.4514
0.35704 | -77.4823
0.36076 | -76.906
0.36448 | -76.9367
0.3682 | -76.7573
0.37192 | -76.3806
0.37564 | -76.3532
0.37936 | -76.841
0.38308 | -76.2216
0.3868 | -76.1533
0.39052 | -77.1702
0.39424 | -78.0972
0.39795 | -77.5635
0.40167 | -76.9916
0.40539 | -78.4111
0.40911 | -78.8465
0.41283 | -77.8676
0.41655 | -77.2089
0.42027 | -78.1211
0.42399 | -78.565
0.42771 | -78.4665
0.43143 | -78.4921
0.43515 | -78.5419
0.43887 | -77.5022
0.44259 | -78.5032
0.4463 | -80.1171
0.45002 | -79.3051
0.45374 | -77.956
0.45746 | -77.7857
0.46118 | -78.1729
0.4649 | -78.286
0.46862 | -79.288
0.47234 | -79.0117
0.47606 | -78.2499
0.47978 | -78.1561
0.4835 | -78.7687
0.48722 | -79.983
0.49093 | -80.4653
0.49465 | -80.7571
0.49837 | -80.6862
0.50209 | -80.504
0.50581 | -80.1652
0.50953 | -80.8072
0.51325 | -80.3906
0.51697 | -79.1253
0.52069 | -78.7017
0.52441 | -80.295
0.52813 | -79.8624
0.53185 | -80.1625
0.53556 | -80.3774
0.53928 | -81.015
0.543 | -81.3697
0.54672 | -81.1773
0.55044 | -80.7955
0.55416 | -80.4771
0.55788 | -80.3737
0.5616 | -80.5659
0.56532 | -80.3202
0.56904 | -80.1173
0.57276 | -80.4072
0.57648 | -80.4191
0.5802 | -80.3959
0.58391 | -80.528
0.58763 | -80.9209
0.59135 | -81.4106
0.59507 | -81.5081
0.59879 | -81.0051
0.60251 | -80.737
0.60623 | -80.9359
0.60995 | -81.2601
0.61367 | -81.296
0.61739 | -81.169
0.62111 | -81.2714
0.62483 | -81.7159
0.62855 | -81.971
0.63226 | -81.3985
0.63598 | -80.9466
0.6397 | -81.2282
0.64342 | -81.8978
0.64714 | -82.2431
0.65086 | -81.8138
0.65458 | -81.2673
0.6583 | -81.7269
0.66202 | -82.6041
0.66574 | -82.781
0.66946 | -82.7664
0.67317 | -82.5436
0.67689 | -82.1906
0.68061 | -82.2589
0.68433 | -82.4659
0.68805 | -82.8467
0.69177 | -83.3517
0.69549 | -83.488
0.69921 | -83.3891
0.70293 | -83.5914
0.70665 | -83.5048
0.71037 | -82.9021
0.71409 | -82.6621
0.71781 | -82.8569
0.72152 | -82.7622
0.72524 | -82.3915
0.72896 | -82.5113
0.73268 | -83.0188
0.7364 | -83.1676
0.74012 | -83.2413
0.74384 | -83.1968
0.74756 | -82.8677
0.75128 | -82.5373
0.755 | -82.7259
0.75872 | -83.391
0.76244 | -83.5136
0.76616 | -83.3775
0.76987 | -83.7647
0.77359 | -84.2439
0.77731 | -84.2951
0.78103 | -84.3282
0.78475 | -84.3905
0.78847 | -84.0372
0.79219 | -83.7536
0.79591 | -83.895
0.79963 | -83.951
0.80335 | -83.8595
0.80707 | -83.9359
0.81079 | -84.6257
0.81451 | -85.1832
0.81822 | -84.7409
0.82194 | -84.3969
0.82566 | -84.6025
0.82938 | -84.8099
0.8331 | -84.9625
0.83682 | -84.8016
0.84054 | -84.607
0.84426 | -84.7995
0.84798 | -85.4108
0.8517 | -85.4919
0.85542 | -85.3209
0.85913 | -85.557
0.86285 | -85.4132
0.86657 | -84.8152
0.87029 | -84.7798
0.87401 | -85.2634
0.87773 | -85.1881
0.88145 | -84.561
0.88517 | -84.2103
0.88889 | -84.1273
0.89261 | -84.5572
0.89633 | -85.4975
0.90005 | -85.9815
0.90377 | -85.5317
0.90748 | -85.2609
0.9112 | -85.4764
0.91492 | -85.1426
0.91864 | -84.9937
0.92236 | -85.4287
0.92608 | -85.7961
0.9298 | -85.8227
0.93352 | -85.8104
0.93724 | -85.8063
0.94096 | -85.5516
0.94468 | -84.8422
0.9484 | -84.4204
0.95212 | -84.9467
0.95583 | -85.5287
0.95955 | -85.5458
0.96327 | -85.5795
0.96699 | -85.5062
0.97071 | -85.7282
0.97443 | -86.2263
0.97815 | -86.2385
0.98187 | -86.0872
0.98559 | -86.0391
0.98931 | -86.1529
0.99303 | -86.736
0.99675 | -87.1292
1.0005 | -86.9135
1.0042 | -86.0602
1.0079 | -85.3474
1.0116 | -85.4753
1.0153 | -86.0888
1.0191 | -86.6455
1.0228 | -87.1024
1.0265 | -87.5205
1.0302 | -87.6619
1.0339 | -87.3732
1.0377 | -86.8311
1.0414 | -86.4838
1.0451 | -86.451
1.0488 | -86.6736
1.0525 | -87.0553
1.0563 | -86.8699
1.06 | -86.3383
1.0637 | -86.3942
1.0674 | -86.7667
1.0711 | -86.8715
1.0748 | -87.1122
1.0786 | -87.3802
1.0823 | -87.4273
1.086 | -87.1691
1.0897 | -86.8976
1.0934 | -86.8226
1.0972 | -86.8497
1.1009 | -86.9187
1.1046 | -87.0564
1.1083 | -87.1945
1.112 | -87.3092
1.1158 | -87.4087
1.1195 | -87.4604
1.1232 | -87.4744
1.1269 | -87.4872
1.1306 | -87.5388
1.1344 | -87.5927
1.1381 | -87.4972
1.1418 | -87.2017
1.1455 | -86.9214
1.1492 | -86.869
1.153 | -86.9829
1.1567 | -87.0779
1.1604 | -87.0912
1.1641 | -87.0659
1.1678 | -87.1046
1.1715 | -87.2727
1.1753 | -87.5196
1.179 | -87.685
1.1827 | -87.656
1.1864 | -87.4542
1.1901 | -87.1249
1.1939 | -86.8129
1.1976 | -86.721
1.2013 | -86.9103
1.205 | -87.3369
1.2087 | -87.8277
1.2125 | -88.0829
1.2162 | -88.0458
1.2199 | -87.942
1.2236 | -87.9056
1.2273 | -87.7818
1.2311 | -87.5664
1.2348 | -87.5769
1.2385 | -87.9223
1.2422 | -88.2968
1.2459 | -88.3405
1.2497 | -88.078
1.2534 | -87.8449
1.2571 | -87.7882
1.2608 | -87.8776
1.2645 | -88.0487
1.2682 | -88.0253
1.272 | -87.5904
1.2757 | -87.1447
1.2794 | -87.0905
1.2831 | -87.2737
1.2868 | -87.2714
1.2906 | -87.0009
1.2943 | -86.8762
1.298 | -87.1804
1.3017 | -87.6872
1.3054 | -88.0495
1.3092 | -88.324
1.3129 | -88.6824
1.3166 | -89.0802
1.3203 | -89.3845
1.324 | -89.5982
1.3278 | -89.727
1.3315 | -89.6651
1.3352 | -89.4864
1.3389 | -89.3915
1.3426 | -89.3602
1.3464 | -89.2315
1.3501 | -89.0284
1.3538 | -88.9561
1.3575 | -89.0657
1.3612 | -89.1538
1.3649 | -89.1109
1.3687 | -89.0028
1.3724 | -88.8409
1.3761 | -88.6411
1.3798 | -88.5116
1.3835 | -88.478
1.3873 | -88.4501
1.391 | -88.3847
1.3947 | -88.4567
1.3984 | -88.8594
1.4021 | -89.5294
1.4059 | -90.0966
1.4096 | -90.1648
1.4133 | -89.7156
1.417 | -89.1529
1.4207 | -88.8937
1.4245 | -88.9676
1.4282 | -89.0666
1.4319 | -89.0664
1.4356 | -89.0408
1.4393 | -89.1341
1.443 | -89.4939
1.4468 | -89.872
1.4505 | -89.8123
1.4542 | -89.4686
1.4579 | -89.2681
1.4616 | -89.2177
1.4654 | -89.1845
1.4691 | -89.2261
1.4728 | -89.3082
1.4765 | -89.2478
1.4802 | -89.0615
1.484 | -88.8976
1.4877 | -88.8389
1.4914 | -88.9676
1.4951 | -89.3156
1.4988 | -89.7268
1.5026 | -89.9895
1.5063 | -90.1118
1.51 | -90.2542
1.5137 | -90.4221
1.5174 | -90.4617
1.5212 | -90.334
1.5249 | -90.1586
1.5286 | -90.0185
1.5323 | -89.8987
1.536 | -89.7685
1.5397 | -89.6259
1.5435 | -89.5108
1.5472 | -89.4442
1.5509 | -89.4269
1.5546 | -89.5096
1.5583 | -89.6525
1.5621 | -89.595
1.5658 | -89.1765
1.5695 | -88.6897
1.5732 | -88.6125
1.5769 | -89.0926
1.5807 | -89.7952
1.5844 | -90.1691
1.5881 | -90.0114
1.5918 | -89.7227
1.5955 | -89.6395
1.5993 | -89.6882
1.603 | -89.7396
1.6067 | -89.7676
1.6104 | -89.7848
1.6141 | -89.7673
1.6179 | -89.7142
1.6216 | -89.6559
1.6253 | -89.6569
1.629 | -89.7122
1.6327 | -89.7498
1.6364 | -89.7003
1.6402 | -89.633
1.6439 | -89.6457
1.6476 | -89.7717
1.6513 | -89.9833
1.655 | -90.2049
1.6588 | -90.3821
1.6625 | -90.5153
1.6662 | -90.6026
1.6699 | -90.6081
1.6736 | -90.4464
1.6774 | -90.1673
1.6811 | -89.8838
1.6848 | -89.6982
1.6885 | -89.6778
1.6922 | -89.85
1.696 | -90.1447
1.6997 | -90.4137
1.7034 | -90.5751
1.7071 | -90.6905
1.7108 | -90.8178
1.7146 | -90.9837
1.7183 | -91.1643
1.722 | -91.2577
1.7257 | -91.1343
1.7294 | -90.8096
1.7331 | -90.4577
1.7369 | -90.2602
1.7406 | -90.2999
1.7443 | -90.5485
1.748 | -90.8535
1.7517 | -91.0312
1.7555 | -91.0554
1.7592 | -91.0331
1.7629 | -91.0473
1.7666 | -91.0826
1.7703 | -91.0727
1.7741 | -90.9796
1.7778 | -90.8237
1.7815 | -90.7307
1.7852 | -90.7947
1.7889 | -90.9781
1.7927 | -91.1862
1.7964 | -91.3826
1.8001 | -91.5602
1.8038 | -91.6983
1.8075 | -91.7773
1.8113 | -91.8165
1.815 | -91.8156
1.8187 | -91.7526
1.8224 | -91.6274
1.8261 | -91.5208
1.8298 | -91.5051
1.8336 | -91.546
1.8373 | -91.5486
1.841 | -91.5023
1.8447 | -91.4848
1.8484 | -91.5319
1.8522 | -91.554
1.8559 | -91.4366
1.8596 | -91.2181
1.8633 | -91.0708
1.867 | -91.1149
1.8708 | -91.2851
1.8745 | -91.4499
1.8782 | -91.519
1.8819 | -91.5152
1.8856 | -91.461
1.8894 | -91.3533
1.8931 | -91.2419
1.8968 | -91.215
1.9005 | -91.2993
1.9042 | -91.4621
1.908 | -91.63
1.9117 | -91.7521
1.9154 | -91.7699
1.9191 | -91.6934
1.9228 | -91.572
1.9265 | -91.4993
1.9303 | -91.5395
1.934 | -91.6957
1.9377 | -91.8915
1.9414 | -92
1.9451 | -91.9486
1.9489 | -91.8019
1.9526 | -91.6722
1.9563 | -91.6103
1.96 | -91.5747
1.9637 | -91.5175
1.9675 | -91.3875
1.9712 | -91.2191
1.9749 | -91.1225
1.9786 | -91.2066
1.9823 | -91.448
1.9861 | -91.6968
1.9898 | -91.7822
1.9935 | -91.7014
1.9972 | -91.5947
2.0009 | -91.5813
2.0046 | -91.6666
2.0084 | -91.8014
2.0121 | -91.9053
2.0158 | -91.9513
2.0195 | -91.9528
2.0232 | -91.955
2.027 | -91.9343
2.0307 | -91.8855
2.0344 | -91.8545
2.0381 | -91.9028
2.0418 | -91.9686
2.0456 | -91.9177
2.0493 | -91.7313
2.053 | -91.5649
2.0567 | -91.5416
2.0604 | -91.6753
2.0642 | -91.8898
2.0679 | -92.1042
2.0716 | -92.2669
2.0753 | -92.391
2.079 | -92.524
2.0828 | -92.6698
2.0865 | -92.7688
2.0902 | -92.7893
2.0939 | -92.7636
2.0976 | -92.7343
2.1013 | -92.6988
2.1051 | -92.6132
2.1088 | -92.4383
2.1125 | -92.191
2.1162 | -91.9697
2.1199 | -91.863
2.1237 | -91.8932
2.1274 | -92.0339
2.1311 | -92.1865
2.1348 | -92.2467
2.1385 | -92.2347
2.1423 | -92.2946
2.146 | -92.5686
2.1497 | -92.7931
2.1534 | -92.965
2.1571 | -93.0172
2.1609 | -92.9666
2.1646 | -92.8568
2.1683 | -92.7662
2.172 | -92.7287
2.1757 | -92.7234
2.1795 | -92.7141
2.1832 | -92.6734
2.1869 | -92.6055
2.1906 | -92.5495
2.1943 | -92.5264
2.198 | -92.5355
2.2018 | -92.5744
2.2055 | -92.629
2.2092 | -92.6522
2.2129 | -92.5989
2.2166 | -92.4661
2.2204 | -92.2982
2.2241 | -92.1697
2.2278 | -92.1287
2.2315 | -92.185
2.2352 | -92.3072
2.239 | -92.4502
2.2427 | -92.6176
2.2464 | -92.8454
2.2501 | -93.14
2.2538 | -93.4343
2.2576 | -93.6273
2.2613 | -93.6516
2.265 | -93.5261
2.2687 | -93.3406
2.2724 | -93.1606
2.2762 | -93.0316
2.2799 | -92.9481
2.2836 | -92.8645
2.2873 | -92.7586
2.291 | -92.6579
2.2947 | -92.6222
2.2985 | -92.6755
2.3022 | -92.7797
2.3059 | -92.8525
2.3096 | -92.8689
2.3133 | -92.8747
2.3171 | -92.9233
2.3208 | -93.0275
2.3245 | -93.1662
2.3282 | -93.3094
2.3319 | -93.4104
2.3357 | -93.4713
2.3394 | -93.5149
2.3431 | -93.591
2.3468 | -93.7198
2.3505 | -93.8371
2.3543 | -93.8808
2.358 | -93.8277
2.3617 | -93.7224
2.3654 | -93.6078
2.3691 | -93.5035
2.3729 | -93.433
2.3766 | -93.4183
2.3803 | -93.4803
2.384 | -93.5949
2.3877 | -93.7018
2.3914 | -93.7432
2.3952 | -93.7319
2.3989 | -93.7139
2.4026 | -93.7293
2.4063 | -93.7818
2.41 | -93.8321
2.4138 | -93.8443
2.4175 | -93.8018
2.4212 | -93.7058
2.4249 | -93.5744
2.4286 | -93.444
2.4324 | -93.3383
2.4361 | -93.2873
2.4398 | -93.3191
2.4435 | -93.4112
2.4472 | -93.5082
2.451 | -93.5586
2.4547 | -93.5382
2.4584 | -93.4838
2.4621 | -93.4623
2.4658 | -93.5046
2.4695 | -93.6025
2.4733 | -93.7306
2.477 | -93.8248
2.4807 | -93.8791
2.4844 | -93.9102
2.4881 | -93.935
2.4919 | -93.935
2.4956 | -93.8893
2.4993 | -93.7785
2.503 | -93.6266
2.5067 | -93.4804
2.5105 | -93.3812
2.5142 | -93.3434
2.5179 | -93.3706
2.5216 | -93.4365
2.5253 | -93.5209
2.5291 | -93.6254
2.5328 | -93.7772
2.5365 | -93.9583
2.5402 | -94.0958
2.5439 | -94.1056
2.5477 | -93.9807
2.5514 | -93.7771
2.5551 | -93.6099
2.5588 | -93.5605
2.5625 | -93.6642
2.5662 | -93.8988
2.57 | -94.1995
2.5737 | -94.4799
2.5774 | -94.6649
2.5811 | -94.7032
2.5848 | -94.6042
2.5886 | -94.4416
2.5923 | -94.2919
2.596 | -94.1608
2.5997 | -94.0408
2.6034 | -93.9005
2.6072 | -93.7518
2.6109 | -93.6335
2.6146 | -93.5996
2.6183 | -93.6558
2.622 | -93.773
2.6258 | -93.9209
2.6295 | -94.089
2.6332 | -94.2604
2.6369 | -94.4203
2.6406 | -94.5476
2.6444 | -94.6573
2.6481 | -94.7361
2.6518 | -94.7643
2.6555 | -94.6857
2.6592 | -94.5197
2.6629 | -94.3345
2.6667 | -94.2016
2.6704 | -94.1691
2.6741 | -94.2381
2.6778 | -94.3751
2.6815 | -94.4637
2.6853 | -94.5245
2.689 | -94.5465
2.6927 | -94.5228
2.6964 | -94.4654
2.7001 | -94.4215
2.7039 | -94.391
2.7076 | -94.3633
2.7113 | -94.3294
2.715 | -94.2851
2.7187 | -94.2433
2.7225 | -94.2228
2.7262 | -94.2573
2.7299 | -94.344
2.7336 | -94.4657
2.7373 | -94.6152
2.7411 | -94.7676
2.7448 | -94.8705
2.7485 | -94.8799
2.7522 | -94.811
2.7559 | -94.6926
2.7596 | -94.5757
2.7634 | -94.4962
2.7671 | -94.4869
2.7708 | -94.5531
2.7745 | -94.6754
2.7782 | -94.8198
2.782 | -94.9487
2.7857 | -95.0289
2.7894 | -95.0514
2.7931 | -95.0253
2.7968 | -94.9632
2.8006 | -94.8861
2.8043 | -94.825
2.808 | -94.7956
2.8117 | -94.7979
2.8154 | -94.7963
2.8192 | -94.7739
2.8229 | -94.7285
2.8266 | -94.687
2.8303 | -94.6745
2.834 | -94.7111
2.8378 | -94.8127
2.8415 | -94.9683
2.8452 | -95.1452
2.8489 | -95.2862
2.8526 | -95.3503
2.8563 | -95.3287
2.8601 | -95.2392
2.8638 | -95.1314
2.8675 | -95.0286
2.8712 | -94.9597
2.8749 | -94.9444
2.8787 | -94.9848
2.8824 | -95.0762
2.8861 | -95.1837
2.8898 | -95.2574
2.8935 | -95.2773
2.8973 | -95.2535
2.901 | -95.213
2.9047 | -95.1818
2.9084 | -95.1877
2.9121 | -95.2257
2.9159 | -95.2739
2.9196 | -95.2935
2.9233 | -95.2444
2.927 | -95.1292
2.9307 | -94.9745
2.9344 | -94.838
2.9382 | -94.7702
2.9419 | -94.7819
2.9456 | -94.8604
2.9493 | -94.979
2.953 | -95.1004
2.9568 | -95.1997
2.9605 | -95.2703
2.9642 | -95.3095
2.9679 | -95.2994
2.9716 | -95.2405
2.9754 | -95.1397
2.9791 | -95.0206
2.9828 | -94.9056
2.9865 | -94.8158
2.9902 | -94.7561
2.994 | -94.7391
2.9977 | -94.7641
3.0014 | -94.8222
3.0051 | -94.8922
3.0088 | -94.9406
3.0126 | -94.9764
3.0163 | -95.0163
3.02 | -95.0724
3.0237 | -95.1585
3.0274 | -95.2665
3.0311 | -95.3925
3.0349 | -95.5307
3.0386 | -95.6674
3.0423 | -95.7672
3.046 | -95.8084
3.0497 | -95.7813
3.0535 | -95.7179
3.0572 | -95.6378
3.0609 | -95.5597
3.0646 | -95.5016
3.0683 | -95.4458
3.0721 | -95.3762
3.0758 | -95.3014
3.0795 | -95.2383
3.0832 | -95.2163
3.0869 | -95.2451
3.0907 | -95.3134
3.0944 | -95.3958
3.0981 | -95.4733
3.1018 | -95.5345
3.1055 | -95.595
3.1093 | -95.6533
3.113 | -95.7141
3.1167 | -95.8007
3.1204 | -95.9037
3.1241 | -96.0032
3.1278 | -96.0911
3.1316 | -96.12
3.1353 | -96.0876
3.139 | -95.994
3.1427 | -95.8534
3.1464 | -95.7017
3.1502 | -95.566
3.1539 | -95.4662
3.1576 | -95.425
3.1613 | -95.4462
3.165 | -95.5238
3.1688 | -95.6625
3.1725 | -95.8461
3.1762 | -96.0205
3.1799 | -96.1618
3.1836 | -96.2263
3.1874 | -96.1968
3.1911 | -96.0961
3.1948 | -95.9563
3.1985 | -95.8191
3.2022 | -95.6983
3.206 | -95.5942
3.2097 | -95.5255
3.2134 | -95.4831
3.2171 | -95.4886
3.2208 | -95.4884
3.2245 | -95.5116
3.2283 | -95.5579
3.232 | -95.6133
3.2357 | -95.6792
3.2394 | -95.7634
3.2431 | -95.8635
3.2469 | -95.9923
3.2506 | -96.1436
3.2543 | -96.2919
3.258 | -96.4044
3.2617 | -96.4607
3.2655 | -96.4459
3.2692 | -96.3567
3.2729 | -96.217
3.2766 | -96.0421
3.2803 | -95.8631
3.2841 | -95.7169
3.2878 | -95.6154
3.2915 | -95.5635
3.2952 | -95.5635
3.2989 | -95.6078
3.3026 | -95.6558
3.3064 | -95.6822
3.3101 | -95.6852
3.3138 | -95.6626
3.3175 | -95.6351
3.3212 | -95.6302
3.325 | -95.651
3.3287 | -95.6923
3.3324 | -95.7563
3.3361 | -95.8353
3.3398 | -95.9263
3.3436 | -96.0307
3.3473 | -96.1454
3.351 | -96.2456
3.3547 | -96.3179
3.3584 | -96.3524
3.3622 | -96.355
3.3659 | -96.3299
3.3696 | -96.2824
3.3733 | -96.2231
3.377 | -96.1654
3.3808 | -96.1265
3.3845 | -96.1092
3.3882 | -96.1078
3.3919 | -96.1128
3.3956 | -96.1091
3.3993 | -96.096
3.4031 | -96.0837
3.4068 | -96.1021
3.4105 | -96.1641
3.4142 | -96.2711
3.4179 | -96.4018
3.4217 | -96.5373
3.4254 | -96.6453
3.4291 | -96.7132
3.4328 | -96.7299
3.4365 | -96.7168
3.4403 | -96.6969
3.444 | -96.6891
3.4477 | -96.7039
3.4514 | -96.7307
3.4551 | -96.7446
3.4589 | -96.7339
3.4626 | -96.6853
3.4663 | -96.5973
3.47 | -96.4961
3.4737 | -96.3915
3.4775 | -96.3048
3.4812 | -96.2699
3.4849 | -96.2789
3.4886 | -96.3311
3.4923 | -96.4043
3.496 | -96.4801
3.4998 | -96.5397
3.5035 | -96.566
3.5072 | -96.5637
3.5109 | -96.5555
3.5146 | -96.5631
3.5184 | -96.61
3.5221 | -96.7069
3.5258 | -96.8411
3.5295 | -96.997
3.5332 | -97.1321
3.537 | -97.1917
3.5407 | -97.1665
3.5444 | -97.05
3.5481 | -96.8706
3.5518 | -96.6789
3.5556 | -96.5076
3.5593 | -96.3952
3.563 | -96.3553
3.5667 | -96.3845
3.5704 | -96.473
3.5742 | -96.5982
3.5779 | -96.7118
3.5816 | -96.7713
3.5853 | -96.7753
3.589 | -96.7449
3.5927 | -96.7137
3.5965 | -96.6951
3.6002 | -96.7145
3.6039 | -96.7656
3.6076 | -96.842
3.6113 | -96.928
3.6151 | -96.9935
3.6188 | -97.0332
3.6225 | -97.0354
3.6262 | -97.0083
3.6299 | -96.9634
3.6337 | -96.9197
3.6374 | -96.8918
3.6411 | -96.878
3.6448 | -96.8831
3.6485 | -96.9125
3.6523 | -96.9583
3.656 | -97.0243
3.6597 | -97.1043
3.6634 | -97.1768
3.6671 | -97.2332
3.6709 | -97.2513
3.6746 | -97.2172
3.6783 | -97.1399
3.682 | -97.0281
3.6857 | -96.918
3.6894 | -96.8242
3.6932 | -96.7519
3.6969 | -96.707
3.7006 | -96.6807
3.7043 | -96.6544
3.708 | -96.6186
3.7118 | -96.5695
3.7155 | -96.5322
3.7192 | -96.5285
3.7229 | -96.579
3.7266 | -96.6871
3.7304 | -96.853
3.7341 | -97.0476
3.7378 | -97.2248
3.7415 | -97.3472
3.7452 | -97.3978
3.749 | -97.382
3.7527 | -97.3298
3.7564 | -97.3142
3.7601 | -97.3139
3.7638 | -97.3368
3.7675 | -97.3821
3.7713 | -97.4275
3.775 | -97.4569
3.7787 | -97.4649
3.7824 | -97.4557
3.7861 | -97.4459
3.7899 | -97.4212
3.7936 | -97.3939
3.7973 | -97.3632
3.801 | -97.323
3.8047 | -97.274
3.8085 | -97.2174
3.8122 | -97.1595
3.8159 | -97.0966
3.8196 | -97.0453
3.8233 | -97.0027
3.8271 | -96.9604
3.8308 | -96.9279
3.8345 | -96.9073
3.8382 | -96.8814
3.8419 | -96.8423
3.8457 | -96.7876
3.8494 | -96.726
3.8531 | -96.6771
3.8568 | -96.6428
3.8605 | -96.6374
3.8642 | -96.6665
3.868 | -96.7373
3.8717 | -96.8451
3.8754 | -96.9752
3.8791 | -97.1179
3.8828 | -97.2607
3.8866 | -97.393
3.8903 | -97.5038
3.894 | -97.583
3.8977 | -97.6357
3.9014 | -97.6713
3.9052 | -97.6774
3.9089 | -97.6624
3.9126 | -97.6299
3.9163 | -97.5773
3.92 | -97.5172
3.9238 | -97.4532
3.9275 | -97.3902
3.9312 | -97.3451
3.9349 | -97.3251
3.9386 | -97.3343
3.9424 | -97.3689
3.9461 | -97.4166
3.9498 | -97.4766
3.9535 | -97.5486
3.9572 | -97.6075
3.9609 | -97.6484
3.9647 | -97.6699
3.9684 | -97.6655
3.9721 | -97.6387
3.9758 | -97.6096
3.9795 | -97.5864
3.9833 | -97.5835
3.987 | -97.5907
3.9907 | -97.6022
3.9944 | -97.6217
3.9981 | -97.6428
4.0019 | -97.654
4.0056 | -97.6578
4.0093 | -97.6625
4.013 | -97.6606
4.0167 | -97.6511
4.0205 | -97.6421
4.0242 | -97.6386
4.0279 | -97.6389
4.0316 | -97.6364
4.0353 | -97.6215
4.0391 | -97.6085
4.0428 | -97.6037
4.0465 | -97.6064
4.0502 | -97.6317
4.0539 | -97.6774
4.0576 | -97.7338
4.0614 | -97.7996
4.0651 | -97.8627
4.0688 | -97.8975
4.0725 | -97.8995
4.0762 | -97.8729
4.08 | -97.8197
4.0837 | -97.7615
4.0874 | -97.6995
4.0911 | -97.6551
4.0948 | -97.6365
4.0986 | -97.6325
4.1023 | -97.6377
4.106 | -97.6531
4.1097 | -97.6648
4.1134 | -97.664
4.1172 | -97.6437
4.1209 | -97.609
4.1246 | -97.5636
4.1283 | -97.5193
4.132 | -97.4824
4.1358 | -97.4555
4.1395 | -97.4516
4.1432 | -97.4741
4.1469 | -97.5218
4.1506 | -97.604
4.1543 | -97.7105
4.1581 | -97.8374
4.1618 | -97.9795
4.1655 | -98.1184
4.1692 | -98.2379
4.1729 | -98.3198
4.1767 | -98.3561
4.1804 | -98.3473
4.1841 | -98.2903
4.1878 | -98.1994
4.1915 | -98.0993
4.1953 | -98.0135
4.199 | -97.9545
4.2027 | -97.9356
4.2064 | -97.9558
4.2101 | -97.9975
4.2139 | -98.0493
4.2176 | -98.1046
4.2213 | -98.1465
4.225 | -98.1756
4.2287 | -98.1951
4.2325 | -98.2087
4.2362 | -98.227
4.2399 | -98.2591
4.2436 | -98.3162
4.2473 | -98.4029
4.251 | -98.5094
4.2548 | -98.6102
4.2585 | -98.6887
4.2622 | -98.714
4.2659 | -98.6822
4.2696 | -98.6044
4.2734 | -98.4921
4.2771 | -98.3581
4.2808 | -98.2116
4.2845 | -98.0928
4.2882 | -98.0116
4.292 | -98.0368
4.2957 | -98.0328
4.2994 | -98.0599
4.3031 | -98.1047
4.3068 | -98.1576
4.3106 | -98.2109
4.3143 | -98.2689
4.318 | -98.331
4.3217 | -98.3985
4.3254 | -98.4657
4.3291 | -98.5334
4.3329 | -98.5702
4.3366 | -98.5862
4.3403 | -98.5784
4.344 | -98.5463
4.3477 | -98.4873
4.3515 | -98.4077
4.3552 | -98.3225
4.3589 | -98.2411
4.3626 | -98.172
4.3663 | -98.1155
4.3701 | -98.0785
4.3738 | -98.0638
4.3775 | -98.0708
4.3812 | -98.0972
4.3849 | -98.1362
4.3887 | -98.1808
4.3924 | -98.2221
4.3961 | -98.259
4.3998 | -98.2818
4.4035 | -98.289
4.4073 | -98.2789
4.411 | -98.2638
4.4147 | -98.2545
4.4184 | -98.2763
4.4221 | -98.318
4.4258 | -98.3686
4.4296 | -98.4262
4.4333 | -98.4911
4.437 | -98.5522
4.4407 | -98.5954
4.4444 | -98.6205
4.4482 | -98.6209
4.4519 | -98.6148
4.4556 | -98.6011
4.4593 | -98.5847
4.463 | -98.5625
4.4668 | -98.5361
4.4705 | -98.5121
4.4742 | -98.4975
4.4779 | -98.4758
4.4816 | -98.4588
4.4854 | -98.4466
4.4891 | -98.4299
4.4928 | -98.4172
4.4965 | -98.4181
4.5002 | -98.4294
4.504 | -98.4488
4.5077 | -98.4902
4.5114 | -98.5497
4.5151 | -98.6132
4.5188 | -98.6755
4.5225 | -98.7201
4.5263 | -98.7381
4.53 | -98.7234
4.5337 | -98.6818
4.5374 | -98.6271
4.5411 | -98.562
4.5449 | -98.4891
4.5486 | -98.4198
4.5523 | -98.355
4.556 | -98.2993
4.5597 | -98.2556
4.5635 | -98.2363
4.5672 | -98.2425
4.5709 | -98.2729
4.5746 | -98.327
4.5783 | -98.4038
4.5821 | -98.4996
4.5858 | -98.6088
4.5895 | -98.7211
4.5932 | -98.8274
4.5969 | -98.9225
4.6007 | -98.9903
4.6044 | -99.0119
4.6081 | -98.9851
4.6118 | -98.9276
4.6155 | -98.8607
4.6192 | -98.7974
4.623 | -98.7452
4.6267 | -98.7187
4.6304 | -98.723
4.6341 | -98.7472
4.6378 | -98.7953
4.6416 | -98.8605
4.6453 | -98.93
4.649 | -98.9881
4.6527 | -99.0369
4.6564 | -99.0591
4.6602 | -99.0571
4.6639 | -99.0223
4.6676 | -98.9696
4.6713 | -98.9108
4.675 | -98.8476
4.6788 | -98.7827
4.6825 | -98.7182
4.6862 | -98.6615
4.6899 | -98.6191
4.6936 | -98.5942
4.6974 | -98.5837
4.7011 | -98.582
4.7048 | -98.5847
4.7085 | -98.5821
4.7122 | -98.5755
4.7159 | -98.5628
4.7197 | -98.5425
4.7234 | -98.5163
4.7271 | -98.4879
4.7308 | -98.458
4.7345 | -98.438
4.7383 | -98.4377
4.742 | -98.4652
4.7457 | -98.5239
4.7494 | -98.6096
4.7531 | -98.7164
4.7569 | -98.8348
4.7606 | -98.9552
4.7643 | -99.0701
4.768 | -99.1771
4.7717 | -99.2753
4.7755 | -99.3522
4.7792 | -99.4121
4.7829 | -99.45
4.7866 | -99.4681
4.7903 | -99.4657
4.794 | -99.443
4.7978 | -99.4088
4.8015 | -99.3568
4.8052 | -99.2921
4.8089 | -99.2211
4.8126 | -99.1561
4.8164 | -99.1032
4.8201 | -99.0723
4.8238 | -99.0693
4.8275 | -99.129
4.8312 | -99.1654
4.835 | -99.2104
4.8387 | -99.2572
4.8424 | -99.307
4.8461 | -99.3565
4.8498 | -99.3986
4.8536 | -99.4249
4.8573 | -99.4422
4.861 | -99.4444
4.8647 | -99.4328
4.8684 | -99.4072
4.8722 | -99.3738
4.8759 | -99.3399
4.8796 | -99.3016
4.8833 | -99.2537
4.887 | -99.211
4.8907 | -99.1809
4.8945 | -99.1483
4.8982 | -99.117
4.9019 | -99.0993
4.9056 | -99.1018
4.9093 | -99.1131
4.9131 | -99.1375
4.9168 | -99.1702
4.9205 | -99.2095
4.9242 | -99.2495
4.9279 | -99.2903
4.9317 | -99.3262
4.9354 | -99.3485
4.9391 | -99.3631
4.9428 | -99.3731
4.9465 | -99.3819
4.9503 | -99.3914
4.954 | -99.4089
4.9577 | -99.4344
4.9614 | -99.4697
4.9651 | -99.5038
4.9689 | -99.5428
4.9726 | -99.5849
4.9763 | -99.6223
4.98 | -99.6415
4.9837 | -99.6492
4.9874 | -99.6582
4.9912 | -99.665
4.9949 | -99.6554
4.9986 | -99.642
5.0023 | -99.634
5.006 | -99.6198
5.0098 | -99.5924
5.0135 | -99.5567
5.0172 | -99.5089
5.0209 | -99.4549
5.0246 | -99.3992
5.0284 | -99.3416
5.0321 | -99.2958
5.0358 | -99.276
5.0395 | -99.2803
5.0432 | -99.3082
5.047 | -99.3466
5.0507 | -99.3949
5.0544 | -99.4536
5.0581 | -99.519
5.0618 | -99.5796
5.0656 | -99.6209
5.0693 | -99.6447
5.073 | -99.6633
5.0767 | -99.6648
5.0804 | -99.6533
5.0841 | -99.6526
5.0879 | -99.6589
5.0916 | -99.6657
5.0953 | -99.6771
5.099 | -99.6929
5.1027 | -99.7091
5.1065 | -99.7212
5.1102 | -99.7254
5.1139 | -99.7152
5.1176 | -99.6951
5.1213 | -99.6753
5.1251 | -99.6645
5.1288 | -99.6508
5.1325 | -99.6311
5.1362 | -99.61
5.1399 | -99.5962
5.1437 | -99.5855
5.1474 | -99.5766
5.1511 | -99.568
5.1548 | -99.5618
5.1585 | -99.569
5.1623 | -99.5754
5.166 | -99.5963
5.1697 | -99.6297
5.1734 | -99.6624
5.1771 | -99.6971
5.1808 | -99.7326
5.1846 | -99.7596
5.1883 | -99.7729
5.192 | -99.774
5.1957 | -99.7647
5.1994 | -99.7441
5.2032 | -99.7148
5.2069 | -99.687
5.2106 | -99.6616
5.2143 | -99.6352
5.218 | -99.606
5.2218 | -99.5872
5.2255 | -99.5627
5.2292 | -99.5386
5.2329 | -99.5244
5.2366 | -99.509
5.2404 | -99.4981
5.2441 | -99.4848
5.2478 | -99.4821
5.2515 | -99.4957
5.2552 | -99.5135
5.2589 | -99.5345
5.2627 | -99.5699
5.2664 | -99.615
5.2701 | -99.6661
5.2738 | -99.7214
5.2775 | -99.7805
5.2813 | -99.8446
5.285 | -99.9153
5.2887 | -99.989
5.2924 | -100.0703
5.2961 | -100.1526
5.2999 | -100.2267
5.3036 | -100.2989
5.3073 | -100.3516
5.311 | -100.37
5.3147 | -100.3717
5.3185 | -100.3533
5.3222 | -100.305
5.3259 | -100.2226
5.3296 | -100.1221
5.3333 | -100.0185
5.3371 | -99.9184
5.3408 | -99.8222
5.3445 | -99.7302
5.3482 | -99.6557
5.3519 | -99.5997
5.3556 | -99.5555
5.3594 | -99.5234
5.3631 | -99.5265
5.3668 | -99.5184
5.3705 | -99.5262
5.3742 | -99.5505
5.378 | -99.5872
5.3817 | -99.6289
5.3854 | -99.6823
5.3891 | -99.7451
5.3928 | -99.8072
5.3966 | -99.8564
5.4003 | -99.8984
5.404 | -99.9306
5.4077 | -99.9462
5.4114 | -99.9477
5.4152 | -99.9374
5.4189 | -99.9137
5.4226 | -99.8844
5.4263 | -99.8589
5.43 | -99.8272
5.4338 | -99.8013
5.4375 | -99.7842
5.4412 | -99.7687
5.4449 | -99.759
5.4486 | -99.7594
5.4523 | -99.7692
5.4561 | -99.7807
5.4598 | -99.7948
5.4635 | -99.8214
5.4672 | -99.8598
5.4709 | -99.8891
5.4747 | -99.9204
5.4784 | -99.9536
5.4821 | -99.9773
5.4858 | -99.9889
5.4895 | -99.9932
5.4933 | -99.9968
5.497 | -99.9979
5.5007 | -100.0017
5.5044 | -100.0182
5.5081 | -100.0376
5.5119 | -100.0638
5.5156 | -100.0984
5.5193 | -100.1371
5.523 | -100.1792
5.5267 | -100.2205
5.5305 | -100.2587
5.5342 | -100.2859
5.5379 | -100.3024
5.5416 | -100.3233
5.5453 | -100.342
5.549 | -100.356
5.5528 | -100.3679
5.5565 | -100.3802
5.5602 | -100.3939
5.5639 | -100.4129
5.5676 | -100.4362
5.5714 | -100.4536
5.5751 | -100.4613
5.5788 | -100.4545
5.5825 | -100.4435
5.5862 | -100.4265
5.59 | -100.4107
5.5937 | -100.4046
5.5974 | -100.4067
5.6011 | -100.4207
5.6048 | -100.4399
5.6086 | -100.4707
5.6123 | -100.5231
5.616 | -100.5846
5.6197 | -100.6533
5.6234 | -100.7147
5.6272 | -100.7567
5.6309 | -100.7969
5.6346 | -100.8205
5.6383 | -100.8196
5.642 | -100.8018
5.6457 | -100.7609
5.6495 | -100.7048
5.6532 | -100.6408
5.6569 | -100.5727
5.6606 | -100.5094
5.6643 | -100.4477
5.6681 | -100.3794
5.6718 | -100.3223
5.6755 | -100.278
5.6792 | -100.2434
5.6829 | -100.2162
5.6867 | -100.1995
5.6904 | -100.1917
5.6941 | -100.1888
5.6978 | -100.1853
5.7015 | -100.1883
5.7053 | -100.2055
5.709 | -100.2361
5.7127 | -100.2728
5.7164 | -100.3066
5.7201 | -100.3527
5.7238 | -100.4042
5.7276 | -100.455
5.7313 | -100.4981
5.735 | -100.5425
5.7387 | -100.5896
5.7424 | -100.6415
5.7462 | -100.6908
5.7499 | -100.7477
5.7536 | -100.8062
5.7573 | -100.8467
5.761 | -100.882
5.7648 | -100.9088
5.7685 | -100.9392
5.7722 | -100.9615
5.7759 | -100.9725
5.7796 | -100.9791
5.7834 | -100.9784
5.7871 | -100.9687
5.7908 | -100.9466
5.7945 | -100.9115
5.7982 | -100.8691
5.802 | -100.8193
5.8057 | -100.7627
5.8094 | -100.714
5.8131 | -100.669
5.8168 | -100.6382
5.8205 | -100.6187
5.8243 | -100.6071
5.828 | -100.6023
5.8317 | -100.614
5.8354 | -100.6341
5.8391 | -100.6635
5.8429 | -100.7105
5.8466 | -100.7537
5.8503 | -100.7895
5.854 | -100.8151
5.8577 | -100.8316
5.8615 | -100.8387
5.8652 | -100.8313
5.8689 | -100.8124
5.8726 | -100.7758
5.8763 | -100.7266
5.8801 | -100.6781
5.8838 | -100.6287
5.8875 | -100.5805
5.8912 | -100.5313
5.8949 | -100.4806
5.8987 | -100.4545
5.9024 | -100.4264
5.9061 | -100.3987
5.9098 | -100.3762
5.9135 | -100.3665
5.9172 | -100.3686
5.921 | -100.3754
5.9247 | -100.3939
5.9284 | -100.4251
5.9321 | -100.4678
5.9358 | -100.5182
5.9396 | -100.5812
5.9433 | -100.6503
5.947 | -100.7225
5.9507 | -100.7952
5.9544 | -100.8692
5.9582 | -100.9441
5.9619 | -101.0183
5.9656 | -101.0928
5.9693 | -101.1578
5.973 | -101.2053
5.9768 | -101.2394
5.9805 | -101.2607
5.9842 | -101.2687
5.9879 | -101.251
5.9916 | -101.2134
5.9954 | -101.1676
5.9991 | -101.1173
6.0028 | -101.0636
6.0065 | -101.0136
6.0102 | -100.9677
6.0139 | -100.9278
6.0177 | -100.8898
6.0214 | -100.8515
6.0251 | -100.8124
6.0288 | -100.7759
6.0325 | -100.7469
6.0363 | -100.7201
6.04 | -100.7024
6.0437 | -100.689
6.0474 | -100.6806
6.0511 | -100.6849
6.0549 | -100.7044
6.0586 | -100.7238
6.0623 | -100.7412
6.066 | -100.7582
6.0697 | -100.772
6.0735 | -100.7781
6.0772 | -100.7727
6.0809 | -100.7567
6.0846 | -100.7377
6.0883 | -100.7176
6.0921 | -100.7047
6.0958 | -100.6987
6.0995 | -100.6953
6.1032 | -100.7012
6.1069 | -100.7255
6.1106 | -100.7633
6.1144 | -100.8128
6.1181 | -100.8632
6.1218 | -100.9125
6.1255 | -100.9476
6.1292 | -100.9681
6.133 | -100.9781
6.1367 | -100.9829
6.1404 | -100.9789
6.1441 | -100.9598
6.1478 | -100.9392
6.1516 | -100.9249
6.1553 | -100.9141
6.159 | -100.9062
6.1627 | -100.9056
6.1664 | -100.9163
6.1702 | -100.9368
6.1739 | -100.9657
6.1776 | -100.9994
6.1813 | -101.0332
6.185 | -101.062
6.1887 | -101.0936
6.1925 | -101.1218
6.1962 | -101.1472
6.1999 | -101.1595
6.2036 | -101.165
6.2073 | -101.1706
6.2111 | -101.1792
6.2148 | -101.1935
6.2185 | -101.2163
6.2222 | -101.2465
6.2259 | -101.2765
6.2297 | -101.3126
6.2334 | -101.35
6.2371 | -101.3846
6.2408 | -101.4114
6.2445 | -101.4288
6.2483 | -101.4322
6.252 | -101.4228
6.2557 | -101.399
6.2594 | -101.3701
6.2631 | -101.3365
6.2669 | -101.3068
6.2706 | -101.2737
6.2743 | -101.2455
6.278 | -101.219
6.2817 | -101.1938
6.2854 | -101.1821
6.2892 | -101.1728
6.2929 | -101.1734
6.2966 | -101.1765
6.3003 | -101.1785
6.304 | -101.1816
6.3078 | -101.1848
6.3115 | -101.1891
6.3152 | -101.1959
6.3189 | -101.2077
6.3226 | -101.2155
6.3264 | -101.2329
6.3301 | -101.2541
6.3338 | -101.2713
6.3375 | -101.2909
6.3412 | -101.3075
6.345 | -101.3211
6.3487 | -101.3324
6.3524 | -101.3401
6.3561 | -101.3517
6.3598 | -101.3598
6.3636 | -101.3668
6.3673 | -101.3741
6.371 | -101.3833
6.3747 | -101.3882
6.3784 | -101.3829
6.3821 | -101.3746
6.3859 | -101.3513
6.3896 | -101.3238
6.3933 | -101.2931
6.397 | -101.2572
6.4007 | -101.2137
6.4045 | -101.1618
6.4082 | -101.1177
6.4119 | -101.0762
6.4156 | -101.0441
6.4193 | -101.0141
6.4231 | -100.9998
6.4268 | -100.9954
6.4305 | -100.9962
6.4342 | -101.0053
6.4379 | -101.0026
6.4417 | -101.0006
6.4454 | -100.9943
6.4491 | -100.987
6.4528 | -100.9765
6.4565 | -100.9693
6.4603 | -100.9633
6.464 | -100.9603
6.4677 | -100.963
6.4714 | -100.975
6.4751 | -101.0013
6.4788 | -101.0391
6.4826 | -101.088
6.4863 | -101.1379
6.49 | -101.1955
6.4937 | -101.2528
6.4974 | -101.3058
6.5012 | -101.3581
6.5049 | -101.392
6.5086 | -101.4158
6.5123 | -101.4298
6.516 | -101.4279
6.5198 | -101.4153
6.5235 | -101.3994
6.5272 | -101.3818
6.5309 | -101.3653
6.5346 | -101.353
6.5384 | -101.3387
6.5421 | -101.3353
6.5458 | -101.3326
6.5495 | -101.3444
6.5532 | -101.3604
6.5569 | -101.375
6.5607 | -101.394
6.5644 | -101.4101
6.5681 | -101.4324
6.5718 | -101.4544
6.5755 | -101.469
6.5793 | -101.4837
6.583 | -101.5027
6.5867 | -101.5234
6.5904 | -101.5314
6.5941 | -101.5369
6.5979 | -101.5392
6.6016 | -101.537
6.6053 | -101.5356
6.609 | -101.5295
6.6127 | -101.5201
6.6165 | -101.5177
6.6202 | -101.5123
6.6239 | -101.5063
6.6276 | -101.4989
6.6313 | -101.4937
6.6351 | -101.494
6.6388 | -101.4959
6.6425 | -101.4979
6.6462 | -101.5037
6.6499 | -101.5007
6.6536 | -101.5095
6.6574 | -101.5243
6.6611 | -101.54
6.6648 | -101.5547
6.6685 | -101.5667
6.6722 | -101.5827
6.676 | -101.5996
6.6797 | -101.6156
6.6834 | -101.6295
6.6871 | -101.6329
6.6908 | -101.6373
6.6946 | -101.6404
6.6983 | -101.6409
6.702 | -101.6373
6.7057 | -101.6355
6.7094 | -101.6323
6.7132 | -101.6331
6.7169 | -101.6381
6.7206 | -101.6437
6.7243 | -101.6489
6.728 | -101.6601
6.7318 | -101.673
6.7355 | -101.6888
6.7392 | -101.707
6.7429 | -101.7206
6.7466 | -101.7394
6.7503 | -101.7702
6.7541 | -101.7956
6.7578 | -101.8227
6.7615 | -101.8534
6.7652 | -101.881
6.7689 | -101.9012
6.7727 | -101.9181
6.7764 | -101.9346
6.7801 | -101.9502
6.7838 | -101.9665
6.7875 | -101.9869
6.7913 | -101.9948
6.795 | -102
6.7987 | -102.0143
6.8024 | -102.0255
6.8061 | -102.0335
6.8099 | -102.0471
6.8136 | -102.0532
6.8173 | -102.0472
6.821 | -102.0407
6.8247 | -102.0359
6.8285 | -102.034
6.8322 | -102.0306
6.8359 | -102.0261
6.8396 | -102.027
6.8433 | -102.0373
6.847 | -102.0469
6.8508 | -102.0469
6.8545 | -102.0441
6.8582 | -102.0355
6.8619 | -102.0237
6.8656 | -102.009
6.8694 | -101.9793
6.8731 | -101.9418
6.8768 | -101.9044
6.8805 | -101.8553
6.8842 | -101.8052
6.888 | -101.7708
6.8917 | -101.7453
6.8954 | -101.7241
6.8991 | -101.7012
6.9028 | -101.6874
6.9066 | -101.6886
6.9103 | -101.7017
6.914 | -101.7251
6.9177 | -101.7585
6.9214 | -101.7994
6.9252 | -101.8456
6.9289 | -101.8875
6.9326 | -101.9244
6.9363 | -101.9593
6.94 | -101.9909
6.9437 | -102.0132
6.9475 | -102.033
6.9512 | -102.0509
6.9549 | -102.0741
6.9586 | -102.0931
6.9623 | -102.1041
6.9661 | -102.1133
6.9698 | -102.1244
6.9735 | -102.1244
6.9772 | -102.1253
6.9809 | -102.1231
6.9847 | -102.1217
6.9884 | -102.1174
6.9921 | -102.1072
6.9958 | -102.103
6.9995 | -102.1025
7.0033 | -102.1107
7.007 | -102.1242
7.0107 | -102.1404
7.0144 | -102.1641
7.0181 | -102.1934
7.0218 | -102.2274
7.0256 | -102.2682
7.0293 | -102.3043
7.033 | -102.3407
7.0367 | -102.3846
7.0404 | -102.4251
7.0442 | -102.4593
7.0479 | -102.4887
7.0516 | -102.5101
7.0553 | -102.5275
7.059 | -102.5373
7.0628 | -102.5456
7.0665 | -102.5534
7.0702 | -102.5565
7.0739 | -102.5568
7.0776 | -102.554
7.0814 | -102.5511
7.0851 | -102.5539
7.0888 | -102.548
7.0925 | -102.5424
7.0962 | -102.5339
7.1 | -102.5174
7.1037 | -102.4961
7.1074 | -102.4653
7.1111 | -102.4226
7.1148 | -102.3726
7.1185 | -102.3146
7.1223 | -102.2564
7.126 | -102.2008
7.1297 | -102.1412
7.1334 | -102.0798
7.1371 | -102.0269
7.1409 | -101.9814
7.1446 | -101.9382
7.1483 | -101.9059
7.152 | -101.8791
7.1557 | -101.8618
7.1595 | -101.8532
7.1632 | -101.8521
7.1669 | -101.8602
7.1706 | -101.8789
7.1743 | -101.9063
7.1781 | -101.9446
7.1818 | -101.9872
7.1855 | -102.0306
7.1892 | -102.0779
7.1929 | -102.1264
7.1967 | -102.1762
7.2004 | -102.2235
7.2041 | -102.2629
7.2078 | -102.3019
7.2115 | -102.3411
7.2152 | -102.3807
7.219 | -102.414
7.2227 | -102.4457
7.2264 | -102.4685
7.2301 | -102.4802
7.2338 | -102.4923
7.2376 | -102.5077
7.2413 | -102.5138
7.245 | -102.5211
7.2487 | -102.5236
7.2524 | -102.5199
7.2562 | -102.5087
7.2599 | -102.5031
7.2636 | -102.4997
7.2673 | -102.4947
7.271 | -102.4981
7.2748 | -102.4989
7.2785 | -102.4971
7.2822 | -102.4965
7.2859 | -102.4932
7.2896 | -102.491
7.2934 | -102.4928
7.2971 | -102.4926
7.3008 | -102.4923
7.3045 | -102.4889
7.3082 | -102.4811
7.3119 | -102.4729
7.3157 | -102.4595
7.3194 | -102.4402
7.3231 | -102.4226
7.3268 | -102.4108
7.3305 | -102.3922
7.3343 | -102.377
7.338 | -102.3568
7.3417 | -102.3347
7.3454 | -102.3118
7.3491 | -102.292
7.3529 | -102.2715
7.3566 | -102.2571
7.3603 | -102.2463
7.364 | -102.2339
7.3677 | -102.2189
7.3715 | -102.2036
7.3752 | -102.1964
7.3789 | -102.182
7.3826 | -102.1741
7.3863 | -102.1701
7.3901 | -102.1664
7.3938 | -102.1679
7.3975 | -102.1713
7.4012 | -102.1754
7.4049 | -102.1819
7.4086 | -102.1915
7.4124 | -102.2062
7.4161 | -102.2229
7.4198 | -102.2407
7.4235 | -102.2585
7.4272 | -102.2776
7.431 | -102.3012
7.4347 | -102.3263
7.4384 | -102.3572
7.4421 | -102.387
7.4458 | -102.4243
7.4496 | -102.461
7.4533 | -102.4971
7.457 | -102.5321
7.4607 | -102.5638
7.4644 | -102.5849
7.4682 | -102.6049
7.4719 | -102.62
7.4756 | -102.631
7.4793 | -102.6305
7.483 | -102.6186
7.4867 | -102.606
7.4905 | -102.5798
7.4942 | -102.5488
7.4979 | -102.5125
7.5016 | -102.4715
7.5053 | -102.4311
7.5091 | -102.4026
7.5128 | -102.3659
7.5165 | -102.3339
7.5202 | -102.3088
7.5239 | -102.2934
7.5277 | -102.2825
7.5314 | -102.2765
7.5351 | -102.2734
7.5388 | -102.2784
7.5425 | -102.278
7.5463 | -102.2808
7.55 | -102.2866
7.5537 | -102.2889
7.5574 | -102.284
7.5611 | -102.2808
7.5649 | -102.2813
7.5686 | -102.2795
7.5723 | -102.2752
7.576 | -102.2741
7.5797 | -102.2689
7.5835 | -102.2675
7.5872 | -102.2632
7.5909 | -102.266
7.5946 | -102.2679
7.5983 | -102.2718
7.602 | -102.2841
7.6058 | -102.2978
7.6095 | -102.3137
7.6132 | -102.3376
7.6169 | -102.3653
7.6206 | -102.3929
7.6244 | -102.4258
7.6281 | -102.4633
7.6318 | -102.4976
7.6355 | -102.5298
7.6392 | -102.5578
7.643 | -102.5838
7.6467 | -102.6087
7.6504 | -102.6321
7.6541 | -102.6499
7.6578 | -102.6714
7.6616 | -102.6876
7.6653 | -102.7026
7.669 | -102.7238
7.6727 | -102.7431
7.6764 | -102.7562
7.6801 | -102.766
7.6839 | -102.7792
7.6876 | -102.7912
7.6913 | -102.8019
7.695 | -102.8104
7.6987 | -102.8048
7.7025 | -102.7971
7.7062 | -102.7834
7.7099 | -102.7676
7.7136 | -102.7394
7.7173 | -102.701
7.7211 | -102.6578
7.7248 | -102.6145
7.7285 | -102.5637
7.7322 | -102.5081
7.7359 | -102.4616
7.7397 | -102.4185
7.7434 | -102.3817
7.7471 | -102.3487
7.7508 | -102.3232
7.7545 | -102.3098
7.7583 | -102.3094
7.762 | -102.3113
7.7657 | -102.3188
7.7694 | -102.3385
7.7731 | -102.3639
7.7768 | -102.3949
7.7806 | -102.4247
7.7843 | -102.4581
7.788 | -102.4921
7.7917 | -102.5159
7.7954 | -102.5279
7.7992 | -102.5335
7.8029 | -102.5327
7.8066 | -102.5228
7.8103 | -102.502
7.814 | -102.475
7.8178 | -102.4364
7.8215 | -102.3892
7.8252 | -102.339
7.8289 | -102.2878
7.8326 | -102.2369
7.8364 | -102.1855
7.8401 | -102.1349
7.8438 | -102.0909
7.8475 | -102.0589
7.8512 | -102.0303
7.855 | -102.0099
7.8587 | -101.996
7.8624 | -101.9937
7.8661 | -101.9992
7.8698 | -102.0123
7.8735 | -102.0324
7.8773 | -102.0626
7.881 | -102.0991
7.8847 | -102.1311
7.8884 | -102.1694
7.8921 | -102.2081
7.8959 | -102.2497
7.8996 | -102.2865
7.9033 | -102.3276
7.907 | -102.3698
7.9107 | -102.413
7.9145 | -102.4492
7.9182 | -102.4805
7.9219 | -102.5075
7.9256 | -102.536
7.9293 | -102.567
7.9331 | -102.5976
7.9368 | -102.6248
7.9405 | -102.6473
7.9442 | -102.6751
7.9479 | -102.7001
7.9516 | -102.7246
7.9554 | -102.7504
7.9591 | -102.7726
7.9628 | -102.7921
7.9665 | -102.8025
7.9702 | -102.8087
7.974 | -102.813
7.9777 | -102.81
7.9814 | -102.802
7.9851 | -102.788
7.9888 | -102.7663
7.9926 | -102.7425
7.9963 | -102.7214
8 | -102.7071
8.0037 | -102.6885
8.0074 | -102.6742
8.0112 | -102.6644
8.0149 | -102.6618
8.0186 | -102.6648
8.0223 | -102.6782
8.026 | -102.6902
8.0298 | -102.7127
8.0335 | -102.7418
8.0372 | -102.7729
8.0409 | -102.8056
8.0446 | -102.8363
8.0483 | -102.8547
8.0521 | -102.8782
8.0558 | -102.9032
8.0595 | -102.9245
8.0632 | -102.934
8.0669 | -102.9379
8.0707 | -102.9388
8.0744 | -102.9333
8.0781 | -102.9254
8.0818 | -102.9171
8.0855 | -102.9142
8.0893 | -102.9062
8.093 | -102.9002
8.0967 | -102.896
8.1004 | -102.8932
8.1041 | -102.8974
8.1079 | -102.9079
8.1116 | -102.9214
8.1153 | -102.9306
8.119 | -102.9385
8.1227 | -102.9573
8.1265 | -102.9811
8.1302 | -103.0034
8.1339 | -103.0223
8.1376 | -103.0384
8.1413 | -103.0521
8.145 | -103.0673
8.1488 | -103.0704
8.1525 | -103.0698
8.1562 | -103.0678
8.1599 | -103.0599
8.1636 | -103.0509
8.1674 | -103.0368
8.1711 | -103.0195
8.1748 | -103.004
8.1785 | -102.9879
8.1822 | -102.9789
8.186 | -102.9769
8.1897 | -102.9843
8.1934 | -102.995
8.1971 | -103.0084
8.2008 | -103.0285
8.2046 | -103.0514
8.2083 | -103.0782
8.212 | -103.1146
8.2157 | -103.152
8.2194 | -103.1884
8.2232 | -103.2283
8.2269 | -103.2637
8.2306 | -103.2968
8.2343 | -103.3307
8.238 | -103.3671
8.2417 | -103.4022
8.2455 | -103.4295
8.2492 | -103.4501
8.2529 | -103.4726
8.2566 | -103.4905
8.2603 | -103.5048
8.2641 | -103.5174
8.2678 | -103.5257
8.2715 | -103.5321
8.2752 | -103.5356
8.2789 | -103.5328
8.2827 | -103.5332
8.2864 | -103.5333
8.2901 | -103.5374
8.2938 | -103.5403
8.2975 | -103.536
8.3013 | -103.5305
8.305 | -103.5256
8.3087 | -103.5138
8.3124 | -103.4977
8.3161 | -103.481
8.3199 | -103.4662
8.3236 | -103.4477
8.3273 | -103.4292
8.331 | -103.4068
8.3347 | -103.3771
8.3384 | -103.3469
8.3422 | -103.3174
8.3459 | -103.2814
8.3496 | -103.2468
8.3533 | -103.2143
8.357 | -103.1786
8.3608 | -103.1366
8.3645 | -103.0891
8.3682 | -103.0454
8.3719 | -103.0094
8.3756 | -102.9718
8.3794 | -102.9384
8.3831 | -102.9056
8.3868 | -102.8769
8.3905 | -102.855
8.3942 | -102.8366
8.398 | -102.8222
8.4017 | -102.816
8.4054 | -102.8125
8.4091 | -102.8122
8.4128 | -102.8138
8.4166 | -102.823
8.4203 | -102.8383
8.424 | -102.8617
8.4277 | -102.8882
8.4314 | -102.9248
8.4351 | -102.9626
8.4389 | -102.9978
8.4426 | -103.0326
8.4463 | -103.0636
8.45 | -103.0825
8.4537 | -103.0993
8.4575 | -103.1133
8.4612 | -103.116
8.4649 | -103.1128
8.4686 | -103.1028
8.4723 | -103.0812
8.4761 | -103.052
8.4798 | -103.0179
8.4835 | -102.9882
8.4872 | -102.9601
8.4909 | -102.9285
8.4947 | -102.8948
8.4984 | -102.8701
8.5021 | -102.8524
8.5058 | -102.8419
8.5095 | -102.8358
8.5132 | -102.8409
8.517 | -102.8558
8.5207 | -102.8796
8.5244 | -102.9057
8.5281 | -102.9377
8.5318 | -102.9773
8.5356 | -103.0212
8.5393 | -103.068
8.543 | -103.1137
8.5467 | -103.1585
8.5504 | -103.1997
8.5542 | -103.2339
8.5579 | -103.2659
8.5616 | -103.3021
8.5653 | -103.334
8.569 | -103.3621
8.5728 | -103.3936
8.5765 | -103.4157
8.5802 | -103.4299
8.5839 | -103.4343
8.5876 | -103.4412
8.5914 | -103.4374
8.5951 | -103.4284
8.5988 | -103.4196
8.6025 | -103.4084
8.6062 | -103.3911
8.6099 | -103.3676
8.6137 | -103.3413
8.6174 | -103.3119
8.6211 | -103.2806
8.6248 | -103.2488
8.6285 | -103.2199
8.6323 | -103.1939
8.636 | -103.1707
8.6397 | -103.1495
8.6434 | -103.1324
8.6471 | -103.1195
8.6509 | -103.1187
8.6546 | -103.1206
8.6583 | -103.1235
8.662 | -103.1313
8.6657 | -103.1427
8.6695 | -103.1507
8.6732 | -103.157
8.6769 | -103.1662
8.6806 | -103.1775
8.6843 | -103.1842
8.6881 | -103.1867
8.6918 | -103.189
8.6955 | -103.1907
8.6992 | -103.1851
8.7029 | -103.1761
8.7066 | -103.1684
8.7104 | -103.1573
8.7141 | -103.1475
8.7178 | -103.136
8.7215 | -103.126
8.7252 | -103.1169
8.729 | -103.1035
8.7327 | -103.0903
8.7364 | -103.0791
8.7401 | -103.0677
8.7438 | -103.0609
8.7476 | -103.0598
8.7513 | -103.0643
8.755 | -103.0755
8.7587 | -103.0896
8.7624 | -103.106
8.7662 | -103.1293
8.7699 | -103.1521
8.7736 | -103.1804
8.7773 | -103.2112
8.781 | -103.239
8.7848 | -103.2696
8.7885 | -103.3012
8.7922 | -103.3267
8.7959 | -103.346
8.7996 | -103.3697
8.8033 | -103.3878
8.8071 | -103.4001
8.8108 | -103.4066
8.8145 | -103.4099
8.8182 | -103.4048
8.8219 | -103.3933
8.8257 | -103.3826
8.8294 | -103.372
8.8331 | -103.359
8.8368 | -103.3401
8.8405 | -103.3254
8.8443 | -103.3142
8.848 | -103.3074
8.8517 | -103.3017
8.8554 | -103.2981
8.8591 | -103.3008
8.8629 | -103.3069
8.8666 | -103.3194
8.8703 | -103.3395
8.874 | -103.3572
8.8777 | -103.3761
8.8814 | -103.3962
8.8852 | -103.4085
8.8889 | -103.4236
8.8926 | -103.4386
8.8963 | -103.4454
8.9 | -103.4522
8.9038 | -103.4572
8.9075 | -103.462
8.9112 | -103.4621
8.9149 | -103.452
8.9186 | -103.44
8.9224 | -103.4338
8.9261 | -103.425
8.9298 | -103.4073
8.9335 | -103.3935
8.9372 | -103.3782
8.941 | -103.3624
8.9447 | -103.3459
8.9484 | -103.3377
8.9521 | -103.3301
8.9558 | -103.3223
8.9596 | -103.3167
8.9633 | -103.3149
8.967 | -103.317
8.9707 | -103.3184
8.9744 | -103.319
8.9781 | -103.3238
8.9819 | -103.3323
8.9856 | -103.3418
8.9893 | -103.3509
8.993 | -103.359
8.9967 | -103.3678
9.0005 | -103.3784
9.0042 | -103.3961
9.0079 | -103.4151
9.0116 | -103.4315
9.0153 | -103.4463
9.0191 | -103.4629
9.0228 | -103.4791
9.0265 | -103.4924
9.0302 | -103.4996
9.0339 | -103.4981
9.0377 | -103.4877
9.0414 | -103.4738
9.0451 | -103.4587
9.0488 | -103.4374
9.0525 | -103.4041
9.0563 | -103.3712
9.06 | -103.3379
9.0637 | -103.2979
9.0674 | -103.2627
9.0711 | -103.2308
9.0748 | -103.2023
9.0786 | -103.1739
9.0823 | -103.1543
9.086 | -103.1411
9.0897 | -103.1344
9.0934 | -103.1339
9.0972 | -103.1383
9.1009 | -103.1503
9.1046 | -103.1721
9.1083 | -103.1989
9.112 | -103.2328
9.1158 | -103.274
9.1195 | -103.3225
9.1232 | -103.3657
9.1269 | -103.4133
9.1306 | -103.4595
9.1344 | -103.5016
9.1381 | -103.5376
9.1418 | -103.5722
9.1455 | -103.6078
9.1492 | -103.642
9.153 | -103.666
9.1567 | -103.6842
9.1604 | -103.7
9.1641 | -103.7114
9.1678 | -103.7233
9.1715 | -103.7338
9.1753 | -103.7417
9.179 | -103.7476
9.1827 | -103.7503
9.1864 | -103.7531
9.1901 | -103.7551
9.1939 | -103.7592
9.1976 | -103.7636
9.2013 | -103.7609
9.205 | -103.7525
9.2087 | -103.7432
9.2125 | -103.7357
9.2162 | -103.7303
9.2199 | -103.7257
9.2236 | -103.7224
9.2273 | -103.7241
9.2311 | -103.7266
9.2348 | -103.7258
9.2385 | -103.7268
9.2422 | -103.7325
9.2459 | -103.7382
9.2497 | -103.7462
9.2534 | -103.7598
9.2571 | -103.7746
9.2608 | -103.7886
9.2645 | -103.8029
9.2682 | -103.8185
9.272 | -103.8345
9.2757 | -103.8561
9.2794 | -103.8743
9.2831 | -103.886
9.2868 | -103.8977
9.2906 | -103.9083
9.2943 | -103.917
9.298 | -103.9224
9.3017 | -103.9235
9.3054 | -103.9205
9.3092 | -103.9148
9.3129 | -103.9049
9.3166 | -103.8897
9.3203 | -103.8708
9.324 | -103.8506
9.3278 | -103.8249
9.3315 | -103.7997
9.3352 | -103.7777
9.3389 | -103.758
9.3426 | -103.7431
9.3464 | -103.7293
9.3501 | -103.7144
9.3538 | -103.7016
9.3575 | -103.6934
9.3612 | -103.6889
9.3649 | -103.6837
9.3687 | -103.6866
9.3724 | -103.694
9.3761 | -103.6997
9.3798 | -103.7113
9.3835 | -103.7255
9.3873 | -103.7392
9.391 | -103.7547
9.3947 | -103.7724
9.3984 | -103.7903
9.4021 | -103.8073
9.4059 | -103.8247
9.4096 | -103.8428
9.4133 | -103.8567
9.417 | -103.872
9.4207 | -103.8886
9.4245 | -103.9007
9.4282 | -103.9027
9.4319 | -103.903
9.4356 | -103.9031
9.4393 | -103.9
9.443 | -103.8892
9.4468 | -103.877
9.4505 | -103.8665
9.4542 | -103.8589
9.4579 | -103.8514
9.4616 | -103.8444
9.4654 | -103.8402
9.4691 | -103.8339
9.4728 | -103.8258
9.4765 | -103.8166
9.4802 | -103.8085
9.484 | -103.8033
9.4877 | -103.7974
9.4914 | -103.7931
9.4951 | -103.79
9.4988 | -103.7889
9.5026 | -103.7905
9.5063 | -103.7902
9.51 | -103.7959
9.5137 | -103.8021
9.5174 | -103.8031
9.5212 | -103.8117
9.5249 | -103.8211
9.5286 | -103.8335
9.5323 | -103.8508
9.536 | -103.8684
9.5397 | -103.8836
9.5435 | -103.9018
9.5472 | -103.9203
9.5509 | -103.9368
9.5546 | -103.9533
9.5583 | -103.9688
9.5621 | -103.9838
9.5658 | -104.0045
9.5695 | -104.0231
9.5732 | -104.0372
9.5769 | -104.0487
9.5807 | -104.0587
9.5844 | -104.0689
9.5881 | -104.0764
9.5918 | -104.0775
9.5955 | -104.0748
9.5993 | -104.0721
9.603 | -104.0673
9.6067 | -104.0638
9.6104 | -104.0588
9.6141 | -104.0541
9.6179 | -104.047
9.6216 | -104.0382
9.6253 | -104.03
9.629 | -104.0189
9.6327 | -104.0078
9.6364 | -103.9991
9.6402 | -103.9931
9.6439 | -103.989
9.6476 | -103.9853
9.6513 | -103.987
9.655 | -103.9922
9.6588 | -103.9963
9.6625 | -103.9993
9.6662 | -104.0049
9.6699 | -104.0116
9.6736 | -104.0194
9.6774 | -104.0273
9.6811 | -104.0372
9.6848 | -104.0459
9.6885 | -104.0563
9.6922 | -104.0683
9.696 | -104.0801
9.6997 | -104.0958
9.7034 | -104.1062
9.7071 | -104.111
9.7108 | -104.1147
9.7146 | -104.1182
9.7183 | -104.1221
9.722 | -104.1271
9.7257 | -104.1251
9.7294 | -104.1201
9.7331 | -104.1131
9.7369 | -104.104
9.7406 | -104.0914
9.7443 | -104.0775
9.748 | -104.0635
9.7517 | -104.0524
9.7555 | -104.0354
9.7592 | -104.0178
9.7629 | -103.9987
9.7666 | -103.9801
9.7703 | -103.9643
9.7741 | -103.9489
9.7778 | -103.9326
9.7815 | -103.9143
9.7852 | -103.8975
9.7889 | -103.8797
9.7927 | -103.8659
9.7964 | -103.8571
9.8001 | -103.8497
9.8038 | -103.8422
9.8075 | -103.8374
9.8113 | -103.8386
9.815 | -103.8445
9.8187 | -103.8524
9.8224 | -103.8665
9.8261 | -103.8851
9.8298 | -103.907
9.8336 | -103.9317
9.8373 | -103.9591
9.841 | -103.9908
9.8447 | -104.0276
9.8484 | -104.069
9.8522 | -104.1116
9.8559 | -104.1529
9.8596 | -104.1905
9.8633 | -104.2261
9.867 | -104.2591
9.8708 | -104.2945
9.8745 | -104.3266
9.8782 | -104.357
9.8819 | -104.3904
9.8856 | -104.4178
9.8894 | -104.4407
9.8931 | -104.4634
9.8968 | -104.4839
9.9005 | -104.5085
9.9042 | -104.5272
9.9079 | -104.5441
9.9117 | -104.5596
9.9154 | -104.5698
9.9191 | -104.577
9.9228 | -104.5768
9.9265 | -104.5706
9.9303 | -104.56
9.934 | -104.5517
9.9377 | -104.5367
9.9414 | -104.5165
9.9451 | -104.497
9.9489 | -104.4758
9.9526 | -104.4521
9.9563 | -104.433
9.96 | -104.4132
9.9637 | -104.3966
9.9675 | -104.3841
9.9712 | -104.3758
9.9749 | -104.3695
9.9786 | -104.3639
9.9823 | -104.3635
9.9861 | -104.3643
9.9898 | -104.3638
9.9935 | -104.3682
9.9972 | -104.3734
10.0009 | -104.3772
10.0046 | -104.3823
10.0084 | -104.3869
10.0121 | -104.3896
10.0158 | -104.3912
10.0195 | -104.3926
10.0232 | -104.3881
10.027 | -104.378
10.0307 | -104.3702
10.0344 | -104.3629
10.0381 | -104.3523
10.0418 | -104.3378
10.0456 | -104.3225
10.0493 | -104.3082
10.053 | -104.2921
10.0567 | -104.2772
10.0604 | -104.2577
10.0642 | -104.2387
10.0679 | -104.2244
10.0716 | -104.2151
10.0753 | -104.2065
10.079 | -104.1951
10.0828 | -104.1822
10.0865 | -104.17
10.0902 | -104.1586
10.0939 | -104.1493
10.0976 | -104.1434
10.1013 | -104.1349
10.1051 | -104.1274
10.1088 | -104.1203
10.1125 | -104.113
10.1162 | -104.1067
10.1199 | -104.0997
10.1237 | -104.0965
10.1274 | -104.0899
10.1311 | -104.0846
10.1348 | -104.0821
10.1385 | -104.0785
10.1423 | -104.0754
10.146 | -104.0706
10.1497 | -104.0683
10.1534 | -104.0662
10.1571 | -104.0653
10.1609 | -104.0637
10.1646 | -104.0624
10.1683 | -104.0657
10.172 | -104.0699
10.1757 | -104.0729
10.1795 | -104.0788
10.1832 | -104.0897
10.1869 | -104.1035
10.1906 | -104.121
10.1943 | -104.1509
10.198 | -104.177
10.2018 | -104.2036
10.2055 | -104.2338
10.2092 | -104.2735
10.2129 | -104.3151
10.2166 | -104.3576
10.2204 | -104.3957
10.2241 | -104.4313
10.2278 | -104.4705
10.2315 | -104.5132
10.2352 | -104.5528
10.239 | -104.5884
10.2427 | -104.6201
10.2464 | -104.6445
10.2501 | -104.6606
10.2538 | -104.6719
10.2576 | -104.6765
10.2613 | -104.6778
10.265 | -104.6727
10.2687 | -104.6634
10.2724 | -104.6474
10.2762 | -104.6263
10.2799 | -104.5967
10.2836 | -104.5627
10.2873 | -104.5285
10.291 | -104.4944
10.2947 | -104.4643
10.2985 | -104.4332
10.3022 | -104.4002
10.3059 | -104.3719
10.3096 | -104.3492
10.3133 | -104.3339
10.3171 | -104.3231
10.3208 | -104.3181
10.3245 | -104.3194
10.3282 | -104.3234
10.3319 | -104.3303
10.3357 | -104.3448
10.3394 | -104.3642
10.3431 | -104.3891
10.3468 | -104.4116
10.3505 | -104.4328
10.3543 | -104.4592
10.358 | -104.4802
10.3617 | -104.4951
10.3654 | -104.5116
10.3691 | -104.5255
10.3728 | -104.5368
10.3766 | -104.5458
10.3803 | -104.5482
10.384 | -104.5432
10.3877 | -104.5324
10.3914 | -104.5176
10.3952 | -104.4987
10.3989 | -104.4769
10.4026 | -104.4541
10.4063 | -104.4258
10.41 | -104.398
10.4138 | -104.3689
10.4175 | -104.334
10.4212 | -104.3035
10.4249 | -104.279
10.4286 | -104.2584
10.4324 | -104.242
10.4361 | -104.2294
10.4398 | -104.2195
10.4435 | -104.2126
10.4472 | -104.2101
10.451 | -104.2111
10.4547 | -104.2172
10.4584 | -104.2258
10.4621 | -104.2386
10.4658 | -104.2511
10.4695 | -104.2658
10.4733 | -104.283
10.477 | -104.3037
10.4807 | -104.3274
10.4844 | -104.3515
10.4881 | -104.3797
10.4919 | -104.4075
10.4956 | -104.436
10.4993 | -104.4634
10.503 | -104.4922
10.5067 | -104.5193
10.5105 | -104.5457
10.5142 | -104.5703
10.5179 | -104.5899
10.5216 | -104.6095
10.5253 | -104.6295
10.5291 | -104.6423
10.5328 | -104.6481
10.5365 | -104.6506
10.5402 | -104.6518
10.5439 | -104.6479
10.5477 | -104.6421
10.5514 | -104.634
10.5551 | -104.6227
10.5588 | -104.6101
10.5625 | -104.5952
10.5662 | -104.5776
10.57 | -104.5627
10.5737 | -104.5406
10.5774 | -104.5173
10.5811 | -104.4971
10.5848 | -104.4759
10.5886 | -104.4527
10.5923 | -104.4325
10.596 | -104.4153
10.5997 | -104.3979
10.6034 | -104.3831
10.6072 | -104.3732
10.6109 | -104.3654
10.6146 | -104.3605
10.6183 | -104.362
10.622 | -104.3668
10.6258 | -104.3759
10.6295 | -104.3829
10.6332 | -104.3872
10.6369 | -104.3946
10.6406 | -104.402
10.6444 | -104.4062
10.6481 | -104.4096
10.6518 | -104.4132
10.6555 | -104.4099
10.6592 | -104.4022
10.6629 | -104.3951
10.6667 | -104.3863
10.6704 | -104.374
10.6741 | -104.3543
10.6778 | -104.3333
10.6815 | -104.3177
10.6853 | -104.3076
10.689 | -104.2983
10.6927 | -104.2875
10.6964 | -104.277
10.7001 | -104.2641
10.7039 | -104.2512
10.7076 | -104.2466
10.7113 | -104.2462
10.715 | -104.2465
10.7187 | -104.2491
10.7225 | -104.2574
10.7262 | -104.2709
10.7299 | -104.2934
10.7336 | -104.308
10.7373 | -104.3235
10.7411 | -104.3447
10.7448 | -104.3717
10.7485 | -104.3995
10.7522 | -104.4284
10.7559 | -104.4579
10.7596 | -104.487
10.7634 | -104.5184
10.7671 | -104.5495
10.7708 | -104.5799
10.7745 | -104.6098
10.7782 | -104.6384
10.782 | -104.6643
10.7857 | -104.6931
10.7894 | -104.7211
10.7931 | -104.7439
10.7968 | -104.7569
10.8006 | -104.775
10.8043 | -104.7921
10.808 | -104.8061
10.8117 | -104.8202
10.8154 | -104.8341
10.8192 | -104.8451
10.8229 | -104.8505
10.8266 | -104.8556
10.8303 | -104.8633
10.834 | -104.867
10.8377 | -104.8656
10.8415 | -104.8592
10.8452 | -104.8559
10.8489 | -104.8548
10.8526 | -104.8499
10.8563 | -104.8446
10.8601 | -104.8408
10.8638 | -104.8299
10.8675 | -104.8177
10.8712 | -104.8013
10.8749 | -104.782
10.8787 | -104.7605
10.8824 | -104.7416
10.8861 | -104.7244
10.8898 | -104.711
10.8935 | -104.6935
10.8973 | -104.6747
10.901 | -104.6564
10.9047 | -104.6381
10.9084 | -104.6229
10.9121 | -104.6142
10.9159 | -104.612
10.9196 | -104.6126
10.9233 | -104.6158
10.927 | -104.6247
10.9307 | -104.6353
10.9344 | -104.6476
10.9382 | -104.6631
10.9419 | -104.6805
10.9456 | -104.6999
10.9493 | -104.7246
10.953 | -104.7491
10.9568 | -104.7717
10.9605 | -104.7887
10.9642 | -104.8055
10.9679 | -104.8217
10.9716 | -104.8275
10.9754 | -104.8277
10.9791 | -104.8234
10.9828 | -104.8197
10.9865 | -104.8098
10.9902 | -104.796
10.994 | -104.7817
10.9977 | -104.764
11.0014 | -104.7443
11.0051 | -104.7193
11.0088 | -104.6906
11.0126 | -104.666
11.0163 | -104.6418
11.02 | -104.6171
11.0237 | -104.5975
11.0274 | -104.5806
11.0311 | -104.5708
11.0349 | -104.567
11.0386 | -104.5691
11.0423 | -104.5729
11.046 | -104.5782
11.0497 | -104.5885
11.0535 | -104.6088
11.0572 | -104.6369
11.0609 | -104.6704
11.0646 | -104.7045
11.0683 | -104.7413
11.0721 | -104.781
11.0758 | -104.8197
11.0795 | -104.8623
11.0832 | -104.9015
11.0869 | -104.9352
11.0907 | -104.9725
11.0944 | -105.009
11.0981 | -105.0402
11.1018 | -105.065
11.1055 | -105.0828
11.1093 | -105.0922
11.113 | -105.0976
11.1167 | -105.1032
11.1204 | -105.1012
11.1241 | -105.0918
11.1278 | -105.078
11.1316 | -105.0618
11.1353 | -105.0406
11.139 | -105.0163
11.1427 | -104.9853
11.1464 | -104.9483
11.1502 | -104.9093
11.1539 | -104.8709
11.1576 | -104.8368
11.1613 | -104.8041
11.165 | -104.7714
11.1688 | -104.7442
11.1725 | -104.7215
11.1762 | -104.7047
11.1799 | -104.6891
11.1836 | -104.6769
11.1874 | -104.6686
11.1911 | -104.6621
11.1948 | -104.6554
11.1985 | -104.6482
11.2022 | -104.6473
11.206 | -104.6504
11.2097 | -104.6515
11.2134 | -104.6506
11.2171 | -104.6518
11.2208 | -104.6513
11.2245 | -104.6526
11.2283 | -104.6502
11.232 | -104.6439
11.2357 | -104.638
11.2394 | -104.633
11.2431 | -104.63
11.2469 | -104.6262
11.2506 | -104.6188
11.2543 | -104.6128
11.258 | -104.606
11.2617 | -104.5986
11.2655 | -104.5973
11.2692 | -104.5966
11.2729 | -104.5938
11.2766 | -104.5882
11.2803 | -104.5802
11.2841 | -104.5757
11.2878 | -104.5741
11.2915 | -104.5747
11.2952 | -104.5745
11.2989 | -104.5728
11.3026 | -104.5723
11.3064 | -104.5735
11.3101 | -104.5793
11.3138 | -104.5829
11.3175 | -104.5842
11.3212 | -104.5808
11.325 | -104.5765
11.3287 | -104.5765
11.3324 | -104.5729
11.3361 | -104.5682
11.3398 | -104.5622
11.3436 | -104.5572
11.3473 | -104.555
11.351 | -104.5535
11.3547 | -104.5531
11.3584 | -104.5518
11.3622 | -104.5516
11.3659 | -104.55
11.3696 | -104.5508
11.3733 | -104.5531
11.377 | -104.5565
11.3808 | -104.5585
11.3845 | -104.5608
11.3882 | -104.5641
11.3919 | -104.5751
11.3956 | -104.5852
11.3993 | -104.5957
11.4031 | -104.6051
11.4068 | -104.6138
11.4105 | -104.627
11.4142 | -104.6401
11.4179 | -104.6507
11.4217 | -104.6634
11.4254 | -104.675
11.4291 | -104.6834
11.4328 | -104.6897
11.4365 | -104.6941
11.4403 | -104.6996
11.444 | -104.706
11.4477 | -104.7138
11.4514 | -104.7174
11.4551 | -104.7174
11.4589 | -104.7214
11.4626 | -104.7271
11.4663 | -104.7338
11.47 | -104.7391
11.4737 | -104.7445
11.4775 | -104.749
11.4812 | -104.7579
11.4849 | -104.7678
11.4886 | -104.7812
11.4923 | -104.7957
11.496 | -104.8084
11.4998 | -104.8231
11.5035 | -104.8312
11.5072 | -104.8391
11.5109 | -104.8509
11.5146 | -104.8568
11.5184 | -104.8634
11.5221 | -104.8661
11.5258 | -104.8714
11.5295 | -104.8742
11.5332 | -104.8732
11.537 | -104.8721
11.5407 | -104.8717
11.5444 | -104.8681
11.5481 | -104.8611
11.5518 | -104.8561
11.5556 | -104.8504
11.5593 | -104.8407
11.563 | -104.8319
11.5667 | -104.823
11.5704 | -104.8136
11.5742 | -104.8067
11.5779 | -104.7984
11.5816 | -104.7911
11.5853 | -104.7853
11.589 | -104.7825
11.5927 | -104.7823
11.5965 | -104.7833
11.6002 | -104.787
11.6039 | -104.7927
11.6076 | -104.7996
11.6113 | -104.8072
11.6151 | -104.8152
11.6188 | -104.8231
11.6225 | -104.8338
11.6262 | -104.847
11.6299 | -104.8618
11.6337 | -104.8784
11.6374 | -104.8949
11.6411 | -104.9126
11.6448 | -104.9266
11.6485 | -104.9396
11.6523 | -104.9559
11.656 | -104.9733
11.6597 | -104.9929
11.6634 | -105.0105
11.6671 | -105.0281
11.6709 | -105.0468
11.6746 | -105.0636
11.6783 | -105.0752
11.682 | -105.0859
11.6857 | -105.0936
11.6894 | -105.1045
11.6932 | -105.1118
11.6969 | -105.1148
11.7006 | -105.1141
11.7043 | -105.1095
11.708 | -105.1056
11.7118 | -105.1022
11.7155 | -105.0999
11.7192 | -105.0931
11.7229 | -105.0845
11.7266 | -105.0752
11.7304 | -105.0694
11.7341 | -105.0635
11.7378 | -105.0576
11.7415 | -105.0526
11.7452 | -105.0471
11.749 | -105.0408
11.7527 | -105.0337
11.7564 | -105.0263
11.7601 | -105.0166
11.7638 | -105.0089
11.7675 | -105.0012
11.7713 | -104.9941
11.775 | -104.989
11.7787 | -104.9852
11.7824 | -104.9814
11.7861 | -104.9807
11.7899 | -104.9824
11.7936 | -104.9821
11.7973 | -104.984
11.801 | -104.9918
11.8047 | -104.9932
11.8085 | -104.997
11.8122 | -105.0029
11.8159 | -105.0151
11.8196 | -105.0254
11.8233 | -105.028
11.8271 | -105.03
11.8308 | -105.0327
11.8345 | -105.0383
11.8382 | -105.0419
11.8419 | -105.0495
11.8457 | -105.0547
11.8494 | -105.0552
11.8531 | -105.0566
11.8568 | -105.0598
11.8605 | -105.061
11.8642 | -105.0586
11.868 | -105.0576
11.8717 | -105.0586
11.8754 | -105.0569
11.8791 | -105.057
11.8828 | -105.0575
11.8866 | -105.0573
11.8903 | -105.0574
11.894 | -105.0591
11.8977 | -105.0594
11.9014 | -105.0561
11.9052 | -105.0536
11.9089 | -105.0507
11.9126 | -105.0471
11.9163 | -105.0436
11.92 | -105.0425
11.9238 | -105.0419
11.9275 | -105.0389
11.9312 | -105.0368
11.9349 | -105.0337
11.9386 | -105.0308
11.9424 | -105.0255
11.9461 | -105.0212
11.9498 | -105.0179
11.9535 | -105.0134
11.9572 | -105.0058
11.9609 | -105.0011
11.9647 | -104.9989
11.9684 | -104.997
11.9721 | -104.9917
11.9758 | -104.9895
11.9795 | -104.9903
11.9833 | -104.9897
11.987 | -104.988
11.9907 | -104.9853
11.9944 | -104.9863
11.9981 | -104.9907
12.0019 | -104.9915
12.0056 | -104.9898
12.0093 | -104.9899
12.013 | -104.9928
12.0167 | -104.9926
12.0205 | -104.9963
12.0242 | -105.0067
12.0279 | -105.0153
12.0316 | -105.0206
12.0353 | -105.0248
12.0391 | -105.0335
12.0428 | -105.0435
12.0465 | -105.0542
12.0502 | -105.0685
12.0539 | -105.0802
12.0576 | -105.0925
12.0614 | -105.104
12.0651 | -105.1175
12.0688 | -105.1324
12.0725 | -105.1468
12.0762 | -105.1637
12.08 | -105.1814
12.0837 | -105.1965
12.0874 | -105.2101
12.0911 | -105.2236
12.0948 | -105.2323
12.0986 | -105.2378
12.1023 | -105.2365
12.106 | -105.2358
12.1097 | -105.2303
12.1134 | -105.2222
12.1172 | -105.2112
12.1209 | -105.1966
12.1246 | -105.1805
12.1283 | -105.1634
12.132 | -105.1469
12.1358 | -105.1249
12.1395 | -105.1007
12.1432 | -105.0783
12.1469 | -105.06
12.1506 | -105.0394
12.1543 | -105.0159
12.1581 | -104.9921
12.1618 | -104.972
12.1655 | -104.9504
12.1692 | -104.9298
12.1729 | -104.9147
12.1767 | -104.9045
12.1804 | -104.8956
12.1841 | -104.8862
12.1878 | -104.8822
12.1915 | -104.8807
12.1953 | -104.8774
12.199 | -104.8791
12.2027 | -104.8872
12.2064 | -104.8959
12.2101 | -104.9079
12.2139 | -104.9199
12.2176 | -104.9332
12.2213 | -104.947
12.225 | -104.9604
12.2287 | -104.9733
12.2325 | -104.9874
12.2362 | -105.0017
12.2399 | -105.0146
12.2436 | -105.0232
12.2473 | -105.0308
12.251 | -105.0366
12.2548 | -105.0424
12.2585 | -105.0452
12.2622 | -105.0484
12.2659 | -105.0518
12.2696 | -105.0545
12.2734 | -105.0564
12.2771 | -105.0604
12.2808 | -105.0614
12.2845 | -105.0625
12.2882 | -105.0617
12.292 | -105.0606
12.2957 | -105.056
12.2994 | -105.0513
12.3031 | -105.048
12.3068 | -105.0453
12.3106 | -105.0393
12.3143 | -105.032
12.318 | -105.0248
12.3217 | -105.0157
12.3254 | -105.0064
12.3291 | -104.9994
12.3329 | -104.9936
12.3366 | -104.9875
12.3403 | -104.984
12.344 | -104.9798
12.3477 | -104.9765
12.3515 | -104.9732
12.3552 | -104.9725
12.3589 | -104.974
12.3626 | -104.9757
12.3663 | -104.9757
12.3701 | -104.9819
12.3738 | -104.9893
12.3775 | -104.9908
12.3812 | -104.9968
12.3849 | -105.0089
12.3887 | -105.0204
12.3924 | -105.0318
12.3961 | -105.0437
12.3998 | -105.0556
12.4035 | -105.0673
12.4073 | -105.0838
12.411 | -105.1006
12.4147 | -105.1183
12.4184 | -105.1344
12.4221 | -105.1508
12.4258 | -105.1671
12.4296 | -105.1851
12.4333 | -105.2039
12.437 | -105.2251
12.4407 | -105.2431
12.4444 | -105.2608
12.4482 | -105.2786
12.4519 | -105.2982
12.4556 | -105.3164
12.4593 | -105.3324
12.463 | -105.3483
12.4668 | -105.3635
12.4705 | -105.3753
12.4742 | -105.3846
12.4779 | -105.3882
12.4816 | -105.3916
12.4854 | -105.3929
12.4891 | -105.3949
12.4928 | -105.3952
12.4965 | -105.3931
12.5002 | -105.3908
12.504 | -105.3901
12.5077 | -105.3879
12.5114 | -105.3842
12.5151 | -105.379
12.5188 | -105.3756
12.5225 | -105.3795
12.5263 | -105.3818
12.53 | -105.3796
12.5337 | -105.3782
12.5374 | -105.3763
12.5411 | -105.3793
12.5449 | -105.3867
12.5486 | -105.3959
12.5523 | -105.4039
12.556 | -105.4149
12.5597 | -105.4273
12.5635 | -105.431
12.5672 | -105.4346
12.5709 | -105.439
12.5746 | -105.4406
12.5783 | -105.4473
12.5821 | -105.4544
12.5858 | -105.4625
12.5895 | -105.4648
12.5932 | -105.4648
12.5969 | -105.4666
12.6007 | -105.4654
12.6044 | -105.4586
12.6081 | -105.4521
12.6118 | -105.445
12.6155 | -105.4343
12.6192 | -105.4224
12.623 | -105.4111
12.6267 | -105.3966
12.6304 | -105.382
12.6341 | -105.3656
12.6378 | -105.3497
12.6416 | -105.3335
12.6453 | -105.3172
12.649 | -105.3042
12.6527 | -105.2969
12.6564 | -105.2905
12.6602 | -105.2862
12.6639 | -105.2826
12.6676 | -105.2803
12.6713 | -105.2805
12.675 | -105.2823
12.6788 | -105.2888
12.6825 | -105.2961
12.6862 | -105.3036
12.6899 | -105.3131
12.6936 | -105.3261
12.6974 | -105.3403
12.7011 | -105.3554
12.7048 | -105.3681
12.7085 | -105.3845
12.7122 | -105.4054
12.7159 | -105.4248
12.7197 | -105.4424
12.7234 | -105.4587
12.7271 | -105.476
12.7308 | -105.4918
12.7345 | -105.5043
12.7383 | -105.5154
12.742 | -105.526
12.7457 | -105.5394
12.7494 | -105.5497
12.7531 | -105.5589
12.7569 | -105.5649
12.7606 | -105.5597
12.7643 | -105.5555
12.768 | -105.5514
12.7717 | -105.5459
12.7755 | -105.5339
12.7792 | -105.5164
12.7829 | -105.4999
12.7866 | -105.4858
12.7903 | -105.4657
12.794 | -105.4395
12.7978 | -105.411
12.8015 | -105.3854
12.8052 | -105.3567
12.8089 | -105.3271
12.8126 | -105.2999
12.8164 | -105.2739
12.8201 | -105.2475
12.8238 | -105.2196
12.8275 | -105.1944
12.8312 | -105.1706
12.835 | -105.1475
12.8387 | -105.1304
12.8424 | -105.1138
12.8461 | -105.0964
12.8498 | -105.0828
12.8536 | -105.0734
12.8573 | -105.0656
12.861 | -105.057
12.8647 | -105.0509
12.8684 | -105.0484
12.8722 | -105.0452
12.8759 | -105.0479
12.8796 | -105.046
12.8833 | -105.0472
12.887 | -105.0486
12.8907 | -105.0508
12.8945 | -105.0533
12.8982 | -105.0551
12.9019 | -105.0565
12.9056 | -105.0595
12.9093 | -105.0605
12.9131 | -105.0623
12.9168 | -105.0689
12.9205 | -105.0748
12.9242 | -105.0796
12.9279 | -105.0861
12.9317 | -105.0897
12.9354 | -105.0929
12.9391 | -105.0968
12.9428 | -105.1057
12.9465 | -105.1159
12.9503 | -105.1242
12.954 | -105.1314
12.9577 | -105.1415
12.9614 | -105.1562
12.9651 | -105.1684
12.9689 | -105.179
12.9726 | -105.1897
12.9763 | -105.2007
12.98 | -105.2116
12.9837 | -105.2237
12.9874 | -105.2309
12.9912 | -105.2417
12.9949 | -105.2546
12.9986 | -105.2623
13.0023 | -105.2684
13.006 | -105.2773
13.0098 | -105.2863
13.0135 | -105.2917
13.0172 | -105.2964
13.0209 | -105.3015
13.0246 | -105.3029
13.0284 | -105.3025
13.0321 | -105.3034
13.0358 | -105.3037
13.0395 | -105.303
13.0432 | -105.3021
13.047 | -105.3009
13.0507 | -105.2979
13.0544 | -105.2959
13.0581 | -105.2933
13.0618 | -105.2876
13.0656 | -105.2817
13.0693 | -105.2802
13.073 | -105.2739
13.0767 | -105.2688
13.0804 | -105.2654
13.0841 | -105.2602
13.0879 | -105.256
13.0916 | -105.253
13.0953 | -105.2492
13.099 | -105.244
13.1027 | -105.2379
13.1065 | -105.229
13.1102 | -105.2251
13.1139 | -105.2234
13.1176 | -105.2173
13.1213 | -105.209
13.1251 | -105.1994
13.1288 | -105.1892
13.1325 | -105.1793
13.1362 | -105.1709
13.1399 | -105.1636
13.1437 | -105.1547
13.1474 | -105.1438
13.1511 | -105.135
13.1548 | -105.1264
13.1585 | -105.1163
13.1623 | -105.1057
13.166 | -105.0943
13.1697 | -105.0835
13.1734 | -105.0706
13.1771 | -105.0588
13.1808 | -105.0451
13.1846 | -105.0299
13.1883 | -105.0137
13.192 | -104.9981
13.1957 | -104.9851
13.1994 | -104.974
13.2032 | -104.9621
13.2069 | -104.9485
13.2106 | -104.9357
13.2143 | -104.9263
13.218 | -104.9202
13.2218 | -104.9096
13.2255 | -104.8981
13.2292 | -104.8859
13.2329 | -104.8766
13.2366 | -104.8663
13.2404 | -104.8528
13.2441 | -104.8391
13.2478 | -104.8294
13.2515 | -104.8196
13.2552 | -104.807
13.2589 | -104.7906
13.2627 | -104.7752
13.2664 | -104.7579
13.2701 | -104.738
13.2738 | -104.7181
13.2775 | -104.696
13.2813 | -104.6711
13.285 | -104.6466
13.2887 | -104.625
13.2924 | -104.597
13.2961 | -104.5715
13.2999 | -104.5466
13.3036 | -104.5235
13.3073 | -104.5
13.311 | -104.4756
13.3147 | -104.449
13.3185 | -104.4247
13.3222 | -104.4026
13.3259 | -104.3807
13.3296 | -104.3598
13.3333 | -104.3449
13.3371 | -104.332
13.3408 | -104.3207
13.3445 | -104.3129
13.3482 | -104.3054
13.3519 | -104.2996
13.3556 | -104.3002
13.3594 | -104.3021
13.3631 | -104.3056
13.3668 | -104.3104
13.3705 | -104.3196
13.3742 | -104.3343
13.378 | -104.3529
13.3817 | -104.3738
13.3854 | -104.3945
13.3891 | -104.4184
13.3928 | -104.4471
13.3966 | -104.4798
13.4003 | -104.5119
13.404 | -104.5487
13.4077 | -104.5888
13.4114 | -104.6325
13.4152 | -104.6732
13.4189 | -104.7143
13.4226 | -104.7566
13.4263 | -104.7981
13.43 | -104.8398
13.4338 | -104.8827
13.4375 | -104.9219
13.4412 | -104.9574
13.4449 | -104.9924
13.4486 | -105.0293
13.4523 | -105.0663
13.4561 | -105.0997
13.4598 | -105.1316
13.4635 | -105.1622
13.4672 | -105.1897
13.4709 | -105.2166
13.4747 | -105.2411
13.4784 | -105.2609
13.4821 | -105.2782
13.4858 | -105.3008
13.4895 | -105.3179
13.4933 | -105.3301
13.497 | -105.3384
13.5007 | -105.3499
13.5044 | -105.3622
13.5081 | -105.3734
13.5119 | -105.3773
13.5156 | -105.3772
13.5193 | -105.3788
13.523 | -105.381
13.5267 | -105.3798
13.5305 | -105.3793
13.5342 | -105.372
13.5379 | -105.3654
13.5416 | -105.3615
13.5453 | -105.3563
13.549 | -105.3488
13.5528 | -105.3434
13.5565 | -105.3381
13.5602 | -105.3316
13.5639 | -105.323
13.5676 | -105.314
13.5714 | -105.3065
13.5751 | -105.2973
13.5788 | -105.2904
13.5825 | -105.2833
13.5862 | -105.2777
13.59 | -105.2719
13.5937 | -105.2695
13.5974 | -105.2647
13.6011 | -105.2609
13.6048 | -105.2586
13.6086 | -105.2611
13.6123 | -105.265
13.616 | -105.2667
13.6197 | -105.2702
13.6234 | -105.2713
13.6272 | -105.2769
13.6309 | -105.284
13.6346 | -105.2875
13.6383 | -105.2912
13.642 | -105.2968
13.6457 | -105.2988
13.6495 | -105.3013
13.6532 | -105.3062
13.6569 | -105.3097
13.6606 | -105.3157
13.6643 | -105.3217
13.6681 | -105.3267
13.6718 | -105.3331
13.6755 | -105.3402
13.6792 | -105.3458
13.6829 | -105.3504
13.6867 | -105.3561
13.6904 | -105.3648
13.6941 | -105.3739
13.6978 | -105.3823
13.7015 | -105.3918
13.7053 | -105.4021
13.709 | -105.4106
13.7127 | -105.4161
13.7164 | -105.4206
13.7201 | -105.4251
13.7238 | -105.4301
13.7276 | -105.4361
13.7313 | -105.4424
13.735 | -105.4449
13.7387 | -105.4482
13.7424 | -105.4527
13.7462 | -105.4566
13.7499 | -105.4588
13.7536 | -105.4611
13.7573 | -105.4629
13.761 | -105.4616
13.7648 | -105.4607
13.7685 | -105.4585
13.7722 | -105.457
13.7759 | -105.4497
13.7796 | -105.4415
13.7834 | -105.4323
13.7871 | -105.4222
13.7908 | -105.4106
13.7945 | -105.3949
13.7982 | -105.3753
13.802 | -105.356
13.8057 | -105.338
13.8094 | -105.3171
13.8131 | -105.2935
13.8168 | -105.2699
13.8205 | -105.2499
13.8243 | -105.2231
13.828 | -105.1945
13.8317 | -105.172
13.8354 | -105.1485
13.8391 | -105.1217
13.8429 | -105.0968
13.8466 | -105.0741
13.8503 | -105.0534
13.854 | -105.0331
13.8577 | -105.0078
13.8615 | -104.9852
13.8652 | -104.966
13.8689 | -104.9458
13.8726 | -104.9289
13.8763 | -104.914
13.8801 | -104.8982
13.8838 | -104.8834
13.8875 | -104.8699
13.8912 | -104.8572
13.8949 | -104.8462
13.8987 | -104.8384
13.9024 | -104.8336
13.9061 | -104.8312
13.9098 | -104.831
13.9135 | -104.8332
13.9172 | -104.8371
13.921 | -104.8434
13.9247 | -104.8538
13.9284 | -104.8698
13.9321 | -104.8882
13.9358 | -104.9085
13.9396 | -104.9302
13.9433 | -104.9536
13.947 | -104.9936
13.9507 | -105.0221
13.9544 | -105.0563
13.9582 | -105.0909
13.9619 | -105.1245
13.9656 | -105.1595
13.9693 | -105.1971
13.973 | -105.2326
13.9768 | -105.2687
13.9805 | -105.3041
13.9842 | -105.3369
13.9879 | -105.3677
13.9916 | -105.395
13.9954 | -105.4215
13.9991 | -105.4481
14.0028 | -105.4709
14.0065 | -105.4863
14.0102 | -105.4986
14.0139 | -105.5109
14.0177 | -105.5181
14.0214 | -105.5187
14.0251 | -105.5228
14.0288 | -105.5253
14.0325 | -105.5232
14.0363 | -105.5186
14.04 | -105.5095
14.0437 | -105.5013
14.0474 | -105.4873
14.0511 | -105.4721
14.0549 | -105.4597
14.0586 | -105.45
14.0623 | -105.4405
14.066 | -105.4285
14.0697 | -105.4177
14.0735 | -105.4041
14.0772 | -105.3905
14.0809 | -105.3816
14.0846 | -105.3774
14.0883 | -105.3721
14.0921 | -105.3698
14.0958 | -105.3673
14.0995 | -105.3662
14.1032 | -105.3673
14.1069 | -105.3727
14.1106 | -105.3801
14.1144 | -105.3862
14.1181 | -105.3937
14.1218 | -105.4008
14.1255 | -105.4099
14.1292 | -105.4179
14.133 | -105.4282
14.1367 | -105.4374
14.1404 | -105.4478
14.1441 | -105.4613
14.1478 | -105.4724
14.1516 | -105.4818
14.1553 | -105.4892
14.159 | -105.4904
14.1627 | -105.4943
14.1664 | -105.499
14.1702 | -105.5041
14.1739 | -105.5049
14.1776 | -105.5045
14.1813 | -105.5036
14.185 | -105.5007
14.1887 | -105.4957
14.1925 | -105.4911
14.1962 | -105.4856
14.1999 | -105.4775
14.2036 | -105.4677
14.2073 | -105.4567
14.2111 | -105.4464
14.2148 | -105.4339
14.2185 | -105.4208
14.2222 | -105.4085
14.2259 | -105.3979
14.2297 | -105.3851
14.2334 | -105.3737
14.2371 | -105.3632
14.2408 | -105.354
14.2445 | -105.3455
14.2483 | -105.3363
14.252 | -105.3285
14.2557 | -105.322
14.2594 | -105.3167
14.2631 | -105.3135
14.2669 | -105.3124
14.2706 | -105.3121
14.2743 | -105.3173
14.278 | -105.3255
14.2817 | -105.3346
14.2854 | -105.3439
14.2892 | -105.3545
14.2929 | -105.3689
14.2966 | -105.3847
14.3003 | -105.4026
14.304 | -105.4218
14.3078 | -105.441
14.3115 | -105.4609
14.3152 | -105.4819
14.3189 | -105.5009
14.3226 | -105.5231
14.3264 | -105.5476
14.3301 | -105.5677
14.3338 | -105.5918
14.3375 | -105.6152
14.3412 | -105.6376
14.345 | -105.6573
14.3487 | -105.6768
14.3524 | -105.6963
14.3561 | -105.7183
14.3598 | -105.7408
14.3636 | -105.7631
14.3673 | -105.783
14.371 | -105.803
14.3747 | -105.8194
14.3784 | -105.8395
14.3821 | -105.8607
14.3859 | -105.8803
14.3896 | -105.8991
14.3933 | -105.9167
14.397 | -105.9326
14.4007 | -105.95
14.4045 | -105.9672
14.4082 | -105.9825
14.4119 | -105.9964
14.4156 | -106.0062
14.4193 | -106.017
14.4231 | -106.0245
14.4268 | -106.0289
14.4305 | -106.0303
14.4342 | -106.0278
14.4379 | -106.0228
14.4417 | -106.0185
14.4454 | -106.0148
14.4491 | -106.009
14.4528 | -105.994
14.4565 | -105.9795
14.4603 | -105.9646
14.464 | -105.9446
14.4677 | -105.9223
14.4714 | -105.8965
14.4751 | -105.8695
14.4788 | -105.8373
14.4826 | -105.8051
14.4863 | -105.7697
14.49 | -105.74
14.4937 | -105.7094
14.4974 | -105.674
14.5012 | -105.642
14.5049 | -105.6111
14.5086 | -105.582
14.5123 | -105.5552
14.516 | -105.5309
14.5198 | -105.5077
14.5235 | -105.4863
14.5272 | -105.4705
14.5309 | -105.4535
14.5346 | -105.438
14.5384 | -105.4276
14.5421 | -105.4178
14.5458 | -105.411
14.5495 | -105.4075
14.5532 | -105.4073
14.557 | -105.4079
14.5607 | -105.4108
14.5644 | -105.415
14.5681 | -105.4226
14.5718 | -105.4315
14.5755 | -105.4399
14.5793 | -105.4532
14.583 | -105.4668
14.5867 | -105.4805
14.5904 | -105.4962
14.5941 | -105.5155
14.5979 | -105.5313
14.6016 | -105.551
14.6053 | -105.5718
14.609 | -105.5925
14.6127 | -105.6145
14.6165 | -105.6354
14.6202 | -105.6514
14.6239 | -105.6681
14.6276 | -105.6838
14.6313 | -105.6971
14.6351 | -105.7081
14.6388 | -105.7168
14.6425 | -105.7244
14.6462 | -105.7295
14.6499 | -105.7279
14.6536 | -105.7253
14.6574 | -105.724
14.6611 | -105.7248
14.6648 | -105.7235
14.6685 | -105.7176
14.6722 | -105.7118
14.676 | -105.7052
14.6797 | -105.6963
14.6834 | -105.6907
14.6871 | -105.6866
14.6908 | -105.683
14.6946 | -105.6788
14.6983 | -105.6768
14.702 | -105.6731
14.7057 | -105.6715
14.7094 | -105.6708
14.7132 | -105.6691
14.7169 | -105.6692
14.7206 | -105.6663
14.7243 | -105.6628
14.728 | -105.6608
14.7318 | -105.66
14.7355 | -105.6601
14.7392 | -105.6607
14.7429 | -105.6628
14.7466 | -105.6626
14.7503 | -105.6602
14.7541 | -105.657
14.7578 | -105.6544
14.7615 | -105.6519
14.7652 | -105.6499
14.7689 | -105.6473
14.7727 | -105.6449
14.7764 | -105.6407
14.7801 | -105.6331
14.7838 | -105.6281
14.7875 | -105.6217
14.7913 | -105.6108
14.795 | -105.602
14.7987 | -105.5943
14.8024 | -105.5893
14.8061 | -105.5835
14.8099 | -105.582
14.8136 | -105.5768
14.8173 | -105.5717
14.821 | -105.5683
14.8247 | -105.5651
14.8285 | -105.5644
14.8322 | -105.5686
14.8359 | -105.5727
14.8396 | -105.5743
14.8433 | -105.5783
14.847 | -105.5799
14.8508 | -105.5829
14.8545 | -105.5868
14.8582 | -105.5871
14.8619 | -105.5888
14.8656 | -105.5884
14.8694 | -105.59
14.8731 | -105.5895
14.8768 | -105.5864
14.8805 | -105.5849
14.8842 | -105.5813
14.888 | -105.5778
14.8917 | -105.5759
14.8954 | -105.5679
14.8991 | -105.5569
14.9028 | -105.5459
14.9066 | -105.5329
14.9103 | -105.5223
14.914 | -105.5136
14.9177 | -105.4998
14.9214 | -105.4851
14.9252 | -105.4714
14.9289 | -105.4562
14.9326 | -105.4438
14.9363 | -105.4309
14.94 | -105.4195
14.9437 | -105.408
14.9475 | -105.398
14.9512 | -105.3861
14.9549 | -105.3721
14.9586 | -105.3644
14.9623 | -105.3585
14.9661 | -105.3539
14.9698 | -105.3483
14.9735 | -105.3427
14.9772 | -105.3386
14.9809 | -105.3364
14.9847 | -105.334
14.9884 | -105.3335
14.9921 | -105.3345
14.9958 | -105.3356
14.9995 | -105.3378
15.0033 | -105.3397
15.007 | -105.3406
15.0107 | -105.3412
15.0144 | -105.3402
15.0181 | -105.3384
15.0219 | -105.3265
15.0256 | -105.3226
15.0293 | -105.3165
15.033 | -105.3121
15.0367 | -105.3034
15.0404 | -105.2948
15.0442 | -105.2819
15.0479 | -105.2689
15.0516 | -105.252
15.0553 | -105.2394
15.059 | -105.2245
15.0628 | -105.212
15.0665 | -105.1997
15.0702 | -105.1842
15.0739 | -105.1702
15.0776 | -105.1545
15.0814 | -105.1379
15.0851 | -105.1182
15.0888 | -105.0988
15.0925 | -105.0814
15.0962 | -105.0652
15.1 | -105.0521
15.1037 | -105.038
15.1074 | -105.0219
15.1111 | -105.0102
15.1148 | -104.9999
15.1186 | -104.9873
15.1223 | -104.9774
15.126 | -104.9671
15.1297 | -104.9568
15.1334 | -104.9488
15.1371 | -104.9457
15.1409 | -104.9422
15.1446 | -104.939
15.1483 | -104.9376
15.152 | -104.9378
15.1557 | -104.9392
15.1595 | -104.9448
15.1632 | -104.9512
15.1669 | -104.9601
15.1706 | -104.9716
15.1743 | -104.982
15.1781 | -104.994
15.1818 | -105.0087
15.1855 | -105.025
15.1892 | -105.0434
15.1929 | -105.0596
15.1967 | -105.0766
15.2004 | -105.0977
15.2041 | -105.1189
15.2078 | -105.1392
15.2115 | -105.16
15.2152 | -105.1809
15.219 | -105.2016
15.2227 | -105.2221
15.2264 | -105.2425
15.2301 | -105.2607
15.2338 | -105.2783
15.2376 | -105.2981
15.2413 | -105.3173
15.245 | -105.334
15.2487 | -105.3524
15.2524 | -105.3662
15.2562 | -105.3823
15.2599 | -105.3956
15.2636 | -105.4103
15.2673 | -105.4237
15.271 | -105.4379
15.2748 | -105.4471
15.2785 | -105.4568
15.2822 | -105.4635
15.2859 | -105.4681
15.2896 | -105.4742
15.2934 | -105.4819
15.2971 | -105.4889
15.3008 | -105.4938
15.3045 | -105.4975
15.3082 | -105.4999
15.3119 | -105.499
15.3157 | -105.4965
15.3194 | -105.4912
15.3231 | -105.4843
15.3268 | -105.4789
15.3305 | -105.4695
15.3343 | -105.4547
15.338 | -105.4389
15.3417 | -105.4209
15.3454 | -105.3986
15.3491 | -105.3796
15.3529 | -105.3591
15.3566 | -105.3346
15.3603 | -105.306
15.364 | -105.2743
15.3677 | -105.2454
15.3715 | -105.2177
15.3752 | -105.1867
15.3789 | -105.157
15.3826 | -105.1288
15.3863 | -105.1018
15.3901 | -105.0757
15.3938 | -105.0513
15.3975 | -105.0266
15.4012 | -105.0026
15.4049 | -104.9813
15.4086 | -104.9646
15.4124 | -104.9508
15.4161 | -104.9382
15.4198 | -104.9256
15.4235 | -104.9147
15.4272 | -104.9069
15.431 | -104.9001
15.4347 | -104.8957
15.4384 | -104.8948
15.4421 | -104.8963
15.4458 | -104.8981
15.4496 | -104.9032
15.4533 | -104.9125
15.457 | -104.9212
15.4607 | -104.9315
15.4644 | -104.9396
15.4682 | -104.95
15.4719 | -104.9637
15.4756 | -104.9786
15.4793 | -104.992
15.483 | -105.0063
15.4868 | -105.0186
15.4905 | -105.0314
15.4942 | -105.0433
15.4979 | -105.0556
15.5016 | -105.0689
15.5053 | -105.0815
15.5091 | -105.0926
15.5128 | -105.0983
15.5165 | -105.1038
15.5202 | -105.107
15.5239 | -105.1097
15.5277 | -105.1118
15.5314 | -105.1134
15.5351 | -105.1086
15.5388 | -105.1056
15.5425 | -105.1042
15.5463 | -105.0997
15.55 | -105.0945
15.5537 | -105.0915
15.5574 | -105.088
15.5611 | -105.0846
15.5649 | -105.0813
15.5686 | -105.0785
15.5723 | -105.0747
15.576 | -105.0725
15.5797 | -105.071
15.5834 | -105.0718
15.5872 | -105.0757
15.5909 | -105.0746
15.5946 | -105.0737
15.5983 | -105.0755
15.602 | -105.0758
15.6058 | -105.0781
15.6095 | -105.0833
15.6132 | -105.0904
15.6169 | -105.0972
15.6206 | -105.1059
15.6244 | -105.1134
15.6281 | -105.1202
15.6318 | -105.1275
15.6355 | -105.1354
15.6392 | -105.1454
15.643 | -105.1573
15.6467 | -105.1703
15.6504 | -105.1818
15.6541 | -105.1947
15.6578 | -105.2079
15.6616 | -105.2214
15.6653 | -105.233
15.669 | -105.2473
15.6727 | -105.2606
15.6764 | -105.273
15.6801 | -105.2886
15.6839 | -105.3037
15.6876 | -105.3167
15.6913 | -105.3263
15.695 | -105.3311
15.6987 | -105.3364
15.7025 | -105.3415
15.7062 | -105.345
15.7099 | -105.3461
15.7136 | -105.347
15.7173 | -105.344
15.7211 | -105.3381
15.7248 | -105.3308
15.7285 | -105.3246
15.7322 | -105.3178
15.7359 | -105.3124
15.7397 | -105.3051
15.7434 | -105.2995
15.7471 | -105.293
15.7508 | -105.2846
15.7545 | -105.2733
15.7583 | -105.2665
15.762 | -105.2559
15.7657 | -105.2469
15.7694 | -105.241
15.7731 | -105.2354
15.7768 | -105.2298
15.7806 | -105.2221
15.7843 | -105.2137
15.788 | -105.2103
15.7917 | -105.2071
15.7954 | -105.2061
15.7992 | -105.2038
15.8029 | -105.1999
15.8066 | -105.1973
15.8103 | -105.1962
15.814 | -105.1957
15.8178 | -105.1981
15.8215 | -105.1981
15.8252 | -105.1973
15.8289 | -105.1989
15.8326 | -105.2003
15.8364 | -105.2017
15.8401 | -105.204
15.8438 | -105.2079
15.8475 | -105.2133
15.8512 | -105.221
15.855 | -105.2287
15.8587 | -105.2343
15.8624 | -105.2429
15.8661 | -105.2501
15.8698 | -105.2552
15.8735 | -105.2606
15.8773 | -105.2699
15.881 | -105.278
15.8847 | -105.2874
15.8884 | -105.2961
15.8921 | -105.3061
15.8959 | -105.3135
15.8996 | -105.3195
15.9033 | -105.3264
15.907 | -105.3316
15.9107 | -105.3343
15.9145 | -105.3405
15.9182 | -105.3432
15.9219 | -105.3423
15.9256 | -105.3413
15.9293 | -105.3387
15.9331 | -105.3329
15.9368 | -105.3212
15.9405 | -105.3073
15.9442 | -105.2934
15.9479 | -105.2756
15.9517 | -105.2537
15.9554 | -105.2279
15.9591 | -105.1984
15.9628 | -105.1684
15.9665 | -105.1302
15.9702 | -105.0928
15.974 | -105.0537
15.9777 | -105.0141
15.9814 | -104.9745
15.9851 | -104.9315
15.9888 | -104.8858
15.9926 | -104.839
15.9963 | -104.793
16 | -104.7467
16.0037 | -104.701
16.0074 | -104.6562
16.0112 | -104.6105
16.0149 | -104.5649
16.0186 | -104.5228
16.0223 | -104.4846
16.026 | -104.4472
16.0298 | -104.4168
16.0335 | -104.3881
16.0372 | -104.3619
16.0409 | -104.3388
16.0446 | -104.319
16.0483 | -104.3048
16.0521 | -104.2957
16.0558 | -104.2878
16.0595 | -104.2843
16.0632 | -104.2861
16.0669 | -104.2929
16.0707 | -104.3037
16.0744 | -104.3196
16.0781 | -104.3403
16.0818 | -104.3619
16.0855 | -104.3881
16.0893 | -104.4176
16.093 | -104.4502
16.0967 | -104.4947
16.1004 | -104.5302
16.1041 | -104.5678
16.1079 | -104.6041
16.1116 | -104.641
16.1153 | -104.6768
16.119 | -104.7138
16.1227 | -104.7491
16.1265 | -104.7848
16.1302 | -104.82
16.1339 | -104.8519
16.1376 | -104.8829
16.1413 | -104.9142
16.145 | -104.9381
16.1488 | -104.9634
16.1525 | -104.9876
16.1562 | -105.0022
16.1599 | -105.0137
16.1636 | -105.0216
16.1674 | -105.0307
16.1711 | -105.0371
16.1748 | -105.0396
16.1785 | -105.0404
16.1822 | -105.0371
16.186 | -105.0322
16.1897 | -105.0253
16.1934 | -105.0173
16.1971 | -105.0133
16.2008 | -105.0051
16.2046 | -104.9946
16.2083 | -104.9835
16.212 | -104.9745
16.2157 | -104.9687
16.2194 | -104.9665
16.2232 | -104.9651
16.2269 | -104.9633
16.2306 | -104.9647
16.2343 | -104.9673
16.238 | -104.9748
16.2417 | -104.986
16.2455 | -104.997
16.2492 | -105.0047
16.2529 | -105.0171
16.2566 | -105.0325
16.2603 | -105.049
16.2641 | -105.0675
16.2678 | -105.0884
16.2715 | -105.1106
16.2752 | -105.1334
16.2789 | -105.1576
16.2827 | -105.1824
16.2864 | -105.2089
16.2901 | -105.236
16.2938 | -105.2613
16.2975 | -105.283
16.3013 | -105.3052
16.305 | -105.3252
16.3087 | -105.3441
16.3124 | -105.3659
16.3161 | -105.3849
16.3199 | -105.4009
16.3236 | -105.4166
16.3273 | -105.4321
16.331 | -105.4442
16.3347 | -105.4545
16.3384 | -105.4629
16.3422 | -105.4718
16.3459 | -105.4774
16.3496 | -105.4817
16.3533 | -105.4823
16.357 | -105.4868
16.3608 | -105.4901
16.3645 | -105.4924
16.3682 | -105.4933
16.3719 | -105.4937
16.3756 | -105.4931
16.3794 | -105.4962
16.3831 | -105.4946
16.3868 | -105.4947
16.3905 | -105.4992
16.3942 | -105.5031
16.398 | -105.5078
16.4017 | -105.5113
16.4054 | -105.5105
16.4091 | -105.5132
16.4128 | -105.5158
16.4166 | -105.5211
16.4203 | -105.5268
16.424 | -105.533
16.4277 | -105.5371
16.4314 | -105.5422
16.4351 | -105.5457
16.4389 | -105.5511
16.4426 | -105.5576
16.4463 | -105.5636
16.45 | -105.566
16.4537 | -105.5712
16.4575 | -105.5781
16.4612 | -105.5845
16.4649 | -105.5926
16.4686 | -105.6023
16.4723 | -105.6088
16.4761 | -105.6174
16.4798 | -105.6285
16.4835 | -105.6415
16.4872 | -105.6522
16.4909 | -105.6636
16.4947 | -105.6748
16.4984 | -105.6878
16.5021 | -105.6997
16.5058 | -105.7103
16.5095 | -105.7231
16.5132 | -105.7369
16.517 | -105.7495
16.5207 | -105.7618
16.5244 | -105.7742
16.5281 | -105.7882
16.5318 | -105.8055
16.5356 | -105.8215
16.5393 | -105.8384
16.543 | -105.8546
16.5467 | -105.8701
16.5504 | -105.8867
16.5542 | -105.905
16.5579 | -105.9212
16.5616 | -105.9362
16.5653 | -105.9529
16.569 | -105.972
16.5728 | -105.9919
16.5765 | -106.0079
16.5802 | -106.0263
16.5839 | -106.046
16.5876 | -106.0657
16.5914 | -106.0822
16.5951 | -106.0977
16.5988 | -106.1129
16.6025 | -106.1301
16.6062 | -106.1478
16.6099 | -106.1648
16.6137 | -106.1803
16.6174 | -106.1972
16.6211 | -106.2124
16.6248 | -106.2247
16.6285 | -106.2375
16.6323 | -106.2407
16.636 | -106.2456
16.6397 | -106.2512
16.6434 | -106.259
16.6471 | -106.2632
16.6509 | -106.268
16.6546 | -106.2729
16.6583 | -106.278
16.662 | -106.2818
16.6657 | -106.2826
16.6695 | -106.2799
16.6732 | -106.2758
16.6769 | -106.2721
16.6806 | -106.2679
16.6843 | -106.2596
16.6881 | -106.2522
16.6918 | -106.2409
16.6955 | -106.2233
16.6992 | -106.2079
16.7029 | -106.193
16.7066 | -106.1788
16.7104 | -106.1644
16.7141 | -106.1472
16.7178 | -106.1222
16.7215 | -106.099
16.7252 | -106.0741
16.729 | -106.048
16.7327 | -106.0225
16.7364 | -105.9972
16.7401 | -105.9705
16.7438 | -105.9471
16.7476 | -105.9191
16.7513 | -105.8891
16.755 | -105.8612
16.7587 | -105.8332
16.7624 | -105.8043
16.7662 | -105.7749
16.7699 | -105.7415
16.7736 | -105.7097
16.7773 | -105.6806
16.781 | -105.6501
16.7848 | -105.6222
16.7885 | -105.5951
16.7922 | -105.5665
16.7959 | -105.541
16.7996 | -105.5177
16.8033 | -105.4956
16.8071 | -105.4734
16.8108 | -105.453
16.8145 | -105.4368
16.8182 | -105.4215
16.8219 | -105.407
16.8257 | -105.3947
16.8294 | -105.3839
16.8331 | -105.3757
16.8368 | -105.3661
16.8405 | -105.3558
16.8443 | -105.3431
16.848 | -105.3374
16.8517 | -105.3284
16.8554 | -105.3181
16.8591 | -105.3109
16.8629 | -105.3073
16.8666 | -105.301
16.8703 | -105.2943
16.874 | -105.2891
16.8777 | -105.2839
16.8815 | -105.2763
16.8852 | -105.2673
16.8889 | -105.2573
16.8926 | -105.2443
16.8963 | -105.231
16.9 | -105.2138
16.9038 | -105.1984
16.9075 | -105.1823
16.9112 | -105.1667
16.9149 | -105.1481
16.9186 | -105.127
16.9224 | -105.1048
16.9261 | -105.0837
16.9298 | -105.0627
16.9335 | -105.0403
16.9372 | -105.0179
16.941 | -104.9973
16.9447 | -104.9731
16.9484 | -104.9478
16.9521 | -104.9218
16.9558 | -104.898
16.9596 | -104.8797
16.9633 | -104.8636
16.967 | -104.8465
16.9707 | -104.8293
16.9744 | -104.8107
16.9781 | -104.7914
16.9819 | -104.777
16.9856 | -104.7635
16.9893 | -104.751
16.993 | -104.7419
16.9967 | -104.7346
17.0005 | -104.7287
17.0042 | -104.7249
17.0079 | -104.7222
17.0116 | -104.7233
17.0153 | -104.7261
17.0191 | -104.7296
17.0228 | -104.7355
17.0265 | -104.7419
17.0302 | -104.7484
17.0339 | -104.7577
17.0377 | -104.7683
17.0414 | -104.781
17.0451 | -104.7948
17.0488 | -104.8092
17.0525 | -104.8259
17.0563 | -104.8448
17.06 | -104.8663
17.0637 | -104.8874
17.0674 | -104.909
17.0711 | -104.9329
17.0748 | -104.9594
17.0786 | -104.9863
17.0823 | -105.0113
17.086 | -105.0423
17.0897 | -105.0688
17.0934 | -105.1001
17.0972 | -105.131
17.1009 | -105.1574
17.1046 | -105.1805
17.1083 | -105.2067
17.112 | -105.2321
17.1158 | -105.2562
17.1195 | -105.279
17.1232 | -105.3018
17.1269 | -105.3233
17.1306 | -105.3452
17.1344 | -105.3677
17.1381 | -105.3883
17.1418 | -105.4085
17.1455 | -105.4242
17.1492 | -105.4395
17.153 | -105.4506
17.1567 | -105.4656
17.1604 | -105.4778
17.1641 | -105.4889
17.1678 | -105.4919
17.1715 | -105.4985
17.1753 | -105.5031
17.179 | -105.5058
17.1827 | -105.5108
17.1864 | -105.5136
17.1901 | -105.5175
17.1939 | -105.5201
17.1976 | -105.5238
17.2013 | -105.5232
17.205 | -105.5205
17.2087 | -105.515
17.2125 | -105.5109
17.2162 | -105.5084
17.2199 | -105.5068
17.2236 | -105.5041
17.2273 | -105.4992
17.2311 | -105.4958
17.2348 | -105.4952
17.2385 | -105.4906
17.2422 | -105.4833
17.2459 | -105.4771
17.2497 | -105.4714
17.2534 | -105.4661
17.2571 | -105.463
17.2608 | -105.4578
17.2645 | -105.4519
17.2682 | -105.4456
17.272 | -105.4387
17.2757 | -105.4333
17.2794 | -105.4254
17.2831 | -105.4152
17.2868 | -105.4052
17.2906 | -105.3911
17.2943 | -105.3794
17.298 | -105.3638
17.3017 | -105.3459
17.3054 | -105.3257
17.3092 | -105.3072
17.3129 | -105.2876
17.3166 | -105.2663
17.3203 | -105.2439
17.324 | -105.2243
17.3278 | -105.2045
17.3315 | -105.1801
17.3352 | -105.1549
17.3389 | -105.129
17.3426 | -105.1029
17.3464 | -105.0774
17.3501 | -105.0521
17.3538 | -105.0276
17.3575 | -105.0055
17.3612 | -104.9837
17.3649 | -104.9602
17.3687 | -104.9384
17.3724 | -104.9179
17.3761 | -104.8998
17.3798 | -104.8852
17.3835 | -104.8714
17.3873 | -104.8612
17.391 | -104.8538
17.3947 | -104.8496
17.3984 | -104.85
17.4021 | -104.8491
17.4059 | -104.8501
17.4096 | -104.8562
17.4133 | -104.8632
17.417 | -104.8756
17.4207 | -104.8913
17.4245 | -104.9097
17.4282 | -104.9326
17.4319 | -104.9544
17.4356 | -104.9766
17.4393 | -105.0023
17.443 | -105.0278
17.4468 | -105.0549
17.4505 | -105.0836
17.4542 | -105.1148
17.4579 | -105.1439
17.4616 | -105.1736
17.4654 | -105.2009
17.4691 | -105.2263
17.4728 | -105.2533
17.4765 | -105.2774
17.4802 | -105.2982
17.484 | -105.3165
17.4877 | -105.3363
17.4914 | -105.3549
17.4951 | -105.3697
17.4988 | -105.3811
17.5026 | -105.3898
17.5063 | -105.3937
17.51 | -105.3972
17.5137 | -105.3962
17.5174 | -105.3954
17.5212 | -105.3927
17.5249 | -105.3879
17.5286 | -105.3829
17.5323 | -105.378
17.536 | -105.3674
17.5397 | -105.36
17.5435 | -105.3488
17.5472 | -105.337
17.5509 | -105.3264
17.5546 | -105.3157
17.5583 | -105.3029
17.5621 | -105.292
17.5658 | -105.2795
17.5695 | -105.2624
17.5732 | -105.25
17.5769 | -105.2421
17.5807 | -105.2338
17.5844 | -105.2268
17.5881 | -105.2211
17.5918 | -105.2181
17.5955 | -105.2144
17.5993 | -105.2128
17.603 | -105.2109
17.6067 | -105.2082
17.6104 | -105.2099
17.6141 | -105.2135
17.6179 | -105.2159
17.6216 | -105.2205
17.6253 | -105.2248
17.629 | -105.2295
17.6327 | -105.2354
17.6364 | -105.2418
17.6402 | -105.2496
17.6439 | -105.2569
17.6476 | -105.2635
17.6513 | -105.2725
17.655 | -105.2833
17.6588 | -105.2917
17.6625 | -105.3028
17.6662 | -105.3129
17.6699 | -105.3217
17.6736 | -105.3321
17.6774 | -105.3422
17.6811 | -105.3526
17.6848 | -105.3638
17.6885 | -105.3752
17.6922 | -105.3899
17.696 | -105.4038
17.6997 | -105.4162
17.7034 | -105.4304
17.7071 | -105.4415
17.7108 | -105.4533
17.7146 | -105.4662
17.7183 | -105.4799
17.722 | -105.4933
17.7257 | -105.5046
17.7294 | -105.5161
17.7331 | -105.5272
17.7369 | -105.5393
17.7406 | -105.5508
17.7443 | -105.5624
17.748 | -105.5736
17.7517 | -105.5832
17.7555 | -105.5937
17.7592 | -105.6037
17.7629 | -105.6128
17.7666 | -105.6239
17.7703 | -105.6337
17.7741 | -105.6425
17.7778 | -105.6504
17.7815 | -105.6568
17.7852 | -105.6638
17.7889 | -105.6744
17.7927 | -105.6816
17.7964 | -105.6879
17.8001 | -105.6966
17.8038 | -105.7011
17.8075 | -105.706
17.8113 | -105.7099
17.815 | -105.7121
17.8187 | -105.7139
17.8224 | -105.7191
17.8261 | -105.7177
17.8298 | -105.7148
17.8336 | -105.7127
17.8373 | -105.7105
17.841 | -105.7076
17.8447 | -105.7013
17.8484 | -105.6934
17.8522 | -105.6838
17.8559 | -105.6745
17.8596 | -105.6642
17.8633 | -105.6509
17.867 | -105.6385
17.8708 | -105.6227
17.8745 | -105.6042
17.8782 | -105.5865
17.8819 | -105.5681
17.8856 | -105.5502
17.8894 | -105.5322
17.8931 | -105.5133
17.8968 | -105.4941
17.9005 | -105.4754
17.9042 | -105.4562
17.9079 | -105.4366
17.9117 | -105.4167
17.9154 | -105.3948
17.9191 | -105.3761
17.9228 | -105.3598
17.9265 | -105.3422
17.9303 | -105.3258
17.934 | -105.3114
17.9377 | -105.2984
17.9414 | -105.2865
17.9451 | -105.2759
17.9489 | -105.2674
17.9526 | -105.2624
17.9563 | -105.255
17.96 | -105.2527
17.9637 | -105.2526
17.9675 | -105.2568
17.9712 | -105.262
17.9749 | -105.2694
17.9786 | -105.2778
17.9823 | -105.2903
17.9861 | -105.3032
17.9898 | -105.3162
17.9935 | -105.3308
17.9972 | -105.3464
18.0009 | -105.3637
18.0046 | -105.3853
18.0084 | -105.4081
18.0121 | -105.4308
18.0158 | -105.4551
18.0195 | -105.4808
18.0232 | -105.5056
18.027 | -105.5307
18.0307 | -105.5578
18.0344 | -105.5848
18.0381 | -105.6118
18.0418 | -105.6368
18.0456 | -105.6617
18.0493 | -105.6881
18.053 | -105.7134
18.0567 | -105.7337
18.0604 | -105.7533
18.0642 | -105.7706
18.0679 | -105.7903
18.0716 | -105.8061
18.0753 | -105.8194
18.079 | -105.8295
18.0828 | -105.8369
18.0865 | -105.8423
18.0902 | -105.8468
18.0939 | -105.8485
18.0976 | -105.8494
18.1013 | -105.8509
18.1051 | -105.8493
18.1088 | -105.8463
18.1125 | -105.8394
18.1162 | -105.828
18.1199 | -105.8179
18.1237 | -105.8007
18.1274 | -105.7849
18.1311 | -105.7653
18.1348 | -105.7445
18.1385 | -105.7212
18.1423 | -105.6964
18.146 | -105.6672
18.1497 | -105.639
18.1534 | -105.6063
18.1571 | -105.571
18.1609 | -105.5328
18.1646 | -105.4985
18.1683 | -105.4635
18.172 | -105.4286
18.1757 | -105.391
18.1795 | -105.3534
18.1832 | -105.319
18.1869 | -105.2862
18.1906 | -105.2515
18.1943 | -105.2166
18.198 | -105.1819
18.2018 | -105.1496
18.2055 | -105.1182
18.2092 | -105.0889
18.2129 | -105.0626
18.2166 | -105.0354
18.2204 | -105.0102
18.2241 | -104.9844
18.2278 | -104.9625
18.2315 | -104.9438
18.2352 | -104.9255
18.239 | -104.9081
18.2427 | -104.9094
18.2464 | -104.8989
18.2501 | -104.8932
18.2538 | -104.8891
18.2576 | -104.8869
18.2613 | -104.8869
18.265 | -104.8896
18.2687 | -104.8944
18.2724 | -104.9019
18.2762 | -104.9127
18.2799 | -104.9243
18.2836 | -104.9382
18.2873 | -104.9551
18.291 | -104.9734
18.2947 | -104.9938
18.2985 | -105.0171
18.3022 | -105.0417
18.3059 | -105.0675
18.3096 | -105.0953
18.3133 | -105.1234
18.3171 | -105.1553
18.3208 | -105.19
18.3245 | -105.2239
18.3282 | -105.2554
18.3319 | -105.2893
18.3357 | -105.3234
18.3394 | -105.3544
18.3431 | -105.3873
18.3468 | -105.4184
18.3505 | -105.447
18.3543 | -105.4739
18.358 | -105.4985
18.3617 | -105.5203
18.3654 | -105.5435
18.3691 | -105.5644
18.3728 | -105.5842
18.3766 | -105.603
18.3803 | -105.6166
18.384 | -105.6291
18.3877 | -105.6407
18.3914 | -105.6489
18.3952 | -105.6531
18.3989 | -105.6557
18.4026 | -105.6575
18.4063 | -105.6579
18.41 | -105.6541
18.4138 | -105.651
18.4175 | -105.6467
18.4212 | -105.6412
18.4249 | -105.6301
18.4286 | -105.6194
18.4324 | -105.6073
18.4361 | -105.5948
18.4398 | -105.5789
18.4435 | -105.5622
18.4472 | -105.5437
18.451 | -105.5277
18.4547 | -105.5119
18.4584 | -105.4981
18.4621 | -105.4814
18.4658 | -105.4663
18.4695 | -105.4523
18.4733 | -105.4398
18.477 | -105.4298
18.4807 | -105.42
18.4844 | -105.4122
18.4881 | -105.4042
18.4919 | -105.3972
18.4956 | -105.3936
18.4993 | -105.3901
18.503 | -105.3885
18.5067 | -105.3897
18.5105 | -105.394
18.5142 | -105.3977
18.5179 | -105.403
18.5216 | -105.4102
18.5253 | -105.4203
18.5291 | -105.4286
18.5328 | -105.4409
18.5365 | -105.4535
18.5402 | -105.4663
18.5439 | -105.4825
18.5477 | -105.5018
18.5514 | -105.5224
18.5551 | -105.5435
18.5588 | -105.5631
18.5625 | -105.5839
18.5662 | -105.6054
18.57 | -105.6265
18.5737 | -105.647
18.5774 | -105.6696
18.5811 | -105.6899
18.5848 | -105.7078
18.5886 | -105.7268
18.5923 | -105.7439
18.596 | -105.7603
18.5997 | -105.7787
18.6034 | -105.7983
18.6072 | -105.8124
18.6109 | -105.8213
18.6146 | -105.8299
18.6183 | -105.8383
18.622 | -105.8442
18.6258 | -105.8484
18.6295 | -105.8505
18.6332 | -105.8499
18.6369 | -105.8471
18.6406 | -105.8429
18.6444 | -105.8364
18.6481 | -105.832
18.6518 | -105.8235
18.6555 | -105.8139
18.6592 | -105.8018
18.6629 | -105.7877
18.6667 | -105.7715
18.6704 | -105.7565
18.6741 | -105.7414
18.6778 | -105.7233
18.6815 | -105.7067
18.6853 | -105.6876
18.689 | -105.668
18.6927 | -105.6472
18.6964 | -105.6287
18.7001 | -105.6095
18.7039 | -105.5926
18.7076 | -105.5731
18.7113 | -105.5541
18.715 | -105.5361
18.7187 | -105.5212
18.7225 | -105.5055
18.7262 | -105.493
18.7299 | -105.4804
18.7336 | -105.4685
18.7373 | -105.4575
18.7411 | -105.4476
18.7448 | -105.4416
18.7485 | -105.4376
18.7522 | -105.4343
18.7559 | -105.4316
18.7596 | -105.4323
18.7634 | -105.4354
18.7671 | -105.4388
18.7708 | -105.4442
18.7745 | -105.4518
18.7782 | -105.471
18.782 | -105.4814
18.7857 | -105.4931
18.7894 | -105.506
18.7931 | -105.5192
18.7968 | -105.5371
18.8006 | -105.5537
18.8043 | -105.5718
18.808 | -105.5914
18.8117 | -105.6106
18.8154 | -105.6272
18.8192 | -105.6443
18.8229 | -105.6614
18.8266 | -105.6821
18.8303 | -105.7029
18.834 | -105.7219
18.8377 | -105.7399
18.8415 | -105.754
18.8452 | -105.7667
18.8489 | -105.7805
18.8526 | -105.7953
18.8563 | -105.8085
18.8601 | -105.8202
18.8638 | -105.8294
18.8675 | -105.8362
18.8712 | -105.8432
18.8749 | -105.8519
18.8787 | -105.8595
18.8824 | -105.8627
18.8861 | -105.8652
18.8898 | -105.8676
18.8935 | -105.8716
18.8973 | -105.873
18.901 | -105.8723
18.9047 | -105.8726
18.9084 | -105.8725
18.9121 | -105.8711
18.9159 | -105.8674
18.9196 | -105.8658
18.9233 | -105.8652
18.927 | -105.8633
18.9307 | -105.8609
18.9344 | -105.86
18.9382 | -105.8593
18.9419 | -105.8565
18.9456 | -105.8556
18.9493 | -105.852
18.953 | -105.8524
18.9568 | -105.8502
18.9605 | -105.8495
18.9642 | -105.8491
18.9679 | -105.8521
18.9716 | -105.8558
18.9754 | -105.8609
18.9791 | -105.8667
18.9828 | -105.8727
18.9865 | -105.8769
18.9902 | -105.8822
18.994 | -105.8894
18.9977 | -105.8998
19.0014 | -105.9116
19.0051 | -105.9225
19.0088 | -105.9326
19.0126 | -105.9422
19.0163 | -105.9539
19.02 | -105.9672
19.0237 | -105.9828
19.0274 | -105.9973
19.0311 | -106.0122
19.0349 | -106.0282
19.0386 | -106.0412
19.0423 | -106.0537
19.046 | -106.0693
19.0497 | -106.0825
19.0535 | -106.0968
19.0572 | -106.1108
19.0609 | -106.1253
19.0646 | -106.1395
19.0683 | -106.1492
19.0721 | -106.1575
19.0758 | -106.1646
19.0795 | -106.1715
19.0832 | -106.1774
19.0869 | -106.1854
19.0907 | -106.1922
19.0944 | -106.1977
19.0981 | -106.2014
19.1018 | -106.2024
19.1055 | -106.2042
19.1093 | -106.2028
19.113 | -106.2045
19.1167 | -106.2066
19.1204 | -106.2034
19.1241 | -106.1978
19.1278 | -106.1939
19.1316 | -106.1894
19.1353 | -106.1845
19.139 | -106.1777
19.1427 | -106.1718
19.1464 | -106.1672
19.1502 | -106.1649
19.1539 | -106.1638
19.1576 | -106.1593
19.1613 | -106.1539
19.165 | -106.1465
19.1688 | -106.1397
19.1725 | -106.1345
19.1762 | -106.129
19.1799 | -106.1239
19.1836 | -106.1178
19.1874 | -106.1109
19.1911 | -106.1034
19.1948 | -106.0988
19.1985 | -106.0942
19.2022 | -106.09
19.206 | -106.0827
19.2097 | -106.0719
19.2134 | -106.0643
19.2171 | -106.0573
19.2208 | -106.051
19.2245 | -106.0422
19.2283 | -106.0353
19.232 | -106.0264
19.2357 | -106.0181
19.2394 | -106.0082
19.2431 | -105.9965
19.2469 | -105.9863
19.2506 | -105.9787
19.2543 | -105.9721
19.258 | -105.967
19.2617 | -105.9617
19.2655 | -105.9555
19.2692 | -105.9476
19.2729 | -105.941
19.2766 | -105.9376
19.2803 | -105.9348
19.2841 | -105.9327
19.2878 | -105.9306
19.2915 | -105.9281
19.2952 | -105.9275
19.2989 | -105.9279
19.3026 | -105.9294
19.3064 | -105.9292
19.3101 | -105.9315
19.3138 | -105.9373
19.3175 | -105.9403
19.3212 | -105.9419
19.325 | -105.9455
19.3287 | -105.9496
19.3324 | -105.9539
19.3361 | -105.9571
19.3398 | -105.9608
19.3436 | -105.9663
19.3473 | -105.9723
19.351 | -105.9775
19.3547 | -105.9819
19.3584 | -105.9863
19.3622 | -105.9902
19.3659 | -105.9954
19.3696 | -105.9988
19.3733 | -106.0032
19.377 | -106.0088
19.3808 | -106.0129
19.3845 | -106.02
19.3882 | -106.0261
19.3919 | -106.0285
19.3956 | -106.0338
19.3993 | -106.0411
19.4031 | -106.0464
19.4068 | -106.0503
19.4105 | -106.0522
19.4142 | -106.0556
19.4179 | -106.0557
19.4217 | -106.0553
19.4254 | -106.0569
19.4291 | -106.0594
19.4328 | -106.0622
19.4365 | -106.0642
19.4403 | -106.0648
19.444 | -106.0666
19.4477 | -106.0651
19.4514 | -106.0613
19.4551 | -106.0599
19.4589 | -106.0591
19.4626 | -106.0587
19.4663 | -106.0554
19.47 | -106.0508
19.4737 | -106.047
19.4775 | -106.0436
19.4812 | -106.0408
19.4849 | -106.0401
19.4886 | -106.0371
19.4923 | -106.0325
19.496 | -106.0277
19.4998 | -106.0242
19.5035 | -106.0216
19.5072 | -106.0216
19.5109 | -106.0185
19.5146 | -106.0139
19.5184 | -106.0118
19.5221 | -106.0109
19.5258 | -106.008
19.5295 | -106.0044
19.5332 | -106.0004
19.537 | -106.0005
19.5407 | -106.0014
19.5444 | -106.0025
19.5481 | -106.0027
19.5518 | -106.0022
19.5556 | -106.0038
19.5593 | -106.0072
19.563 | -106.0097
19.5667 | -106.0132
19.5704 | -106.0155
19.5742 | -106.0189
19.5779 | -106.0232
19.5816 | -106.0296
19.5853 | -106.0346
19.589 | -106.0395
19.5927 | -106.0448
19.5965 | -106.0502
19.6002 | -106.0576
19.6039 | -106.0636
19.6076 | -106.0677
19.6113 | -106.0732
19.6151 | -106.0802
19.6188 | -106.0881
19.6225 | -106.094
19.6262 | -106.1001
19.6299 | -106.1064
19.6337 | -106.1114
19.6374 | -106.1141
19.6411 | -106.1163
19.6448 | -106.1194
19.6485 | -106.1235
19.6523 | -106.1248
19.656 | -106.1262
19.6597 | -106.1293
19.6634 | -106.1308
19.6671 | -106.1305
19.6709 | -106.1309
19.6746 | -106.1292
19.6783 | -106.1249
19.682 | -106.122
19.6857 | -106.1214
19.6894 | -106.1188
19.6932 | -106.113
19.6969 | -106.1077
19.7006 | -106.102
19.7043 | -106.0959
19.708 | -106.0897
19.7118 | -106.0845
19.7155 | -106.0786
19.7192 | -106.0732
19.7229 | -106.065
19.7266 | -106.0595
19.7304 | -106.0546
19.7341 | -106.0476
19.7378 | -106.0409
19.7415 | -106.035
19.7452 | -106.0288
19.749 | -106.0217
19.7527 | -106.015
19.7564 | -106.011
19.7601 | -106.0069
19.7638 | -106.0028
19.7675 | -105.9991
19.7713 | -105.9976
19.775 | -105.9953
19.7787 | -105.9919
19.7824 | -105.9886
19.7861 | -105.985
19.7899 | -105.9857
19.7936 | -105.989
19.7973 | -105.9904
19.801 | -105.9932
19.8047 | -105.9983
19.8085 | -106.0055
19.8122 | -106.0125
19.8159 | -106.017
19.8196 | -106.0235
19.8233 | -106.0326
19.8271 | -106.0416
19.8308 | -106.0502
19.8345 | -106.0603
19.8382 | -106.0687
19.8419 | -106.0758
19.8457 | -106.0829
19.8494 | -106.09
19.8531 | -106.0989
19.8568 | -106.1061
19.8605 | -106.1131
19.8642 | -106.1216
19.868 | -106.1302
19.8717 | -106.1388
19.8754 | -106.1457
19.8791 | -106.1518
19.8828 | -106.1565
19.8866 | -106.1605
19.8903 | -106.1661
19.894 | -106.1698
19.8977 | -106.1707
19.9014 | -106.1691
19.9052 | -106.1704
19.9089 | -106.1697
19.9126 | -106.17
19.9163 | -106.1686
19.92 | -106.1676
19.9238 | -106.1627
19.9275 | -106.1578
19.9312 | -106.1528
19.9349 | -106.1465
19.9386 | -106.1382
19.9424 | -106.1295
19.9461 | -106.1225
19.9498 | -106.1148
19.9535 | -106.1056
19.9572 | -106.0942
19.9609 | -106.0852
19.9647 | -106.0733
19.9684 | -106.0593
19.9721 | -106.0444
19.9758 | -106.0291
19.9795 | -106.01
19.9833 | -105.9913
19.987 | -105.9683
19.9907 | -105.9456
19.9944 | -105.9222
19.9981 | -105.9018
20.0019 | -105.8818
20.0056 | -105.8604
20.0093 | -105.8367
20.013 | -105.8108
20.0167 | -105.7861
20.0205 | -105.7601
20.0242 | -105.7299
20.0279 | -105.7019
20.0316 | -105.6741
20.0353 | -105.6458
20.0391 | -105.6136
20.0428 | -105.5822
20.0465 | -105.5514
20.0502 | -105.5197
20.0539 | -105.4875
20.0576 | -105.4565
20.0614 | -105.4275
20.0651 | -105.3973
20.0688 | -105.3665
20.0725 | -105.3359
20.0762 | -105.3033
20.08 | -105.2749
20.0837 | -105.245
20.0874 | -105.2143
20.0911 | -105.1851
20.0948 | -105.1582
20.0986 | -105.1324
20.1023 | -105.1057
20.106 | -105.0788
20.1097 | -105.0553
20.1134 | -105.0352
20.1172 | -105.0153
20.1209 | -104.9965
20.1246 | -104.981
20.1283 | -104.9691
20.132 | -104.96
20.1358 | -104.9508
20.1395 | -104.9428
20.1432 | -104.9393
20.1469 | -104.9365
20.1506 | -104.9354
20.1543 | -104.9374
20.1581 | -104.9411
20.1618 | -104.9481
20.1655 | -104.9577
20.1692 | -104.9696
20.1729 | -104.9839
20.1767 | -104.9986
20.1804 | -105.0157
20.1841 | -105.0353
20.1878 | -105.056
20.1915 | -105.0752
20.1953 | -105.0995
20.199 | -105.122
20.2027 | -105.1456
20.2064 | -105.1703
20.2101 | -105.1981
20.2139 | -105.2226
20.2176 | -105.2469
20.2213 | -105.2697
20.225 | -105.294
20.2287 | -105.3178
20.2325 | -105.3387
20.2362 | -105.3564
20.2399 | -105.3738
20.2436 | -105.3863
20.2473 | -105.3972
20.251 | -105.4055
20.2548 | -105.4126
20.2585 | -105.417
20.2622 | -105.4202
20.2659 | -105.4195
20.2696 | -105.4152
20.2734 | -105.4105
20.2771 | -105.4029
20.2808 | -105.3893
20.2845 | -105.3701
20.2882 | -105.3464
20.292 | -105.3244
20.2957 | -105.2996
20.2994 | -105.2724
20.3031 | -105.2416
20.3068 | -105.2082
20.3106 | -105.1734
20.3143 | -105.1373
20.318 | -105.0979
20.3217 | -105.059
20.3254 | -105.0195
20.3291 | -104.9801
20.3329 | -104.9393
20.3366 | -104.8985
20.3403 | -104.8579
20.344 | -104.8181
20.3477 | -104.7777
20.3515 | -104.7401
20.3552 | -104.7028
20.3589 | -104.6675
20.3626 | -104.6367
20.3663 | -104.6068
20.3701 | -104.5809
20.3738 | -104.5577
20.3775 | -104.5371
20.3812 | -104.5193
20.3849 | -104.5054
20.3887 | -104.5174
20.3924 | -104.5127
20.3961 | -104.5112
20.3998 | -104.5122
20.4035 | -104.5158
20.4073 | -104.5217
20.411 | -104.5315
20.4147 | -104.5452
20.4184 | -104.5596
20.4221 | -104.5756
20.4258 | -104.5954
20.4296 | -104.6188
20.4333 | -104.6425
20.437 | -104.6688
20.4407 | -104.6962
20.4444 | -104.7267
20.4482 | -104.7573
20.4519 | -104.7876
20.4556 | -104.8196
20.4593 | -104.853
20.463 | -104.8856
20.4668 | -104.9189
20.4705 | -104.9529
20.4742 | -104.985
20.4779 | -105.0167
20.4816 | -105.0484
20.4854 | -105.0754
20.4891 | -105.1027
20.4928 | -105.13
20.4965 | -105.1559
20.5002 | -105.1767
20.504 | -105.1951
20.5077 | -105.213
20.5114 | -105.2297
20.5151 | -105.2431
20.5188 | -105.2538
20.5225 | -105.2637
20.5263 | -105.2743
20.53 | -105.2808
20.5337 | -105.2883
20.5374 | -105.2903
20.5411 | -105.2932
20.5449 | -105.2934
20.5486 | -105.2913
20.5523 | -105.2856
20.556 | -105.2819
20.5597 | -105.2747
20.5635 | -105.2667
20.5672 | -105.255
20.5709 | -105.2479
20.5746 | -105.2408
20.5783 | -105.2341
20.5821 | -105.2237
20.5858 | -105.215
20.5895 | -105.2075
20.5932 | -105.1993
20.5969 | -105.193
20.6007 | -105.1897
20.6044 | -105.1841
20.6081 | -105.1802
20.6118 | -105.1756
20.6155 | -105.1727
20.6192 | -105.1685
20.623 | -105.1664
20.6267 | -105.1661
20.6304 | -105.1663
20.6341 | -105.1651
20.6378 | -105.166
20.6416 | -105.1681
20.6453 | -105.1725
20.649 | -105.1752
20.6527 | -105.1787
20.6564 | -105.1828
20.6602 | -105.1873
20.6639 | -105.1931
20.6676 | -105.1995
20.6713 | -105.2062
20.675 | -105.2131
20.6788 | -105.2183
20.6825 | -105.2237
20.6862 | -105.227
20.6899 | -105.2325
20.6936 | -105.2364
20.6974 | -105.2407
20.7011 | -105.2459
20.7048 | -105.252
20.7085 | -105.2574
20.7122 | -105.2635
20.7159 | -105.2692
20.7197 | -105.2755
20.7234 | -105.2805
20.7271 | -105.2852
20.7308 | -105.2895
20.7345 | -105.2942
20.7383 | -105.2951
20.742 | -105.2986
20.7457 | -105.2998
20.7494 | -105.2987
20.7531 | -105.2966
20.7569 | -105.2956
20.7606 | -105.2947
20.7643 | -105.2916
20.768 | -105.2881
20.7717 | -105.2846
20.7755 | -105.2804
20.7792 | -105.2767
20.7829 | -105.27
20.7866 | -105.2639
20.7903 | -105.2578
20.794 | -105.2512
20.7978 | -105.2439
20.8015 | -105.2367
20.8052 | -105.2281
20.8089 | -105.2157
20.8126 | -105.2023
20.8164 | -105.1914
20.8201 | -105.1809
20.8238 | -105.1714
20.8275 | -105.1606
20.8312 | -105.1497
20.835 | -105.1364
20.8387 | -105.1254
20.8424 | -105.1157
20.8461 | -105.1043
20.8498 | -105.0918
20.8536 | -105.0803
20.8573 | -105.0704
20.861 | -105.0616
20.8647 | -105.0478
20.8684 | -105.0348
20.8722 | -105.0234
20.8759 | -105.011
20.8796 | -104.9972
20.8833 | -104.9838
20.887 | -104.9719
20.8907 | -104.9588
20.8945 | -104.945
20.8982 | -104.9329
20.9019 | -104.9202
20.9056 | -104.906
20.9093 | -104.89
20.9131 | -104.8736
20.9168 | -104.855
20.9205 | -104.8362
20.9242 | -104.8117
20.9279 | -104.7929
20.9317 | -104.7765
20.9354 | -104.7557
20.9391 | -104.7337
20.9428 | -104.7112
20.9465 | -104.6876
20.9503 | -104.6625
20.954 | -104.6357
20.9577 | -104.6094
20.9614 | -104.5814
20.9651 | -104.5528
20.9689 | -104.5226
20.9726 | -104.4932
20.9763 | -104.4654
20.98 | -104.4369
20.9837 | -104.4068
20.9874 | -104.3769
20.9912 | -104.3475
20.9949 | -104.317
20.9986 | -104.2887
21.0023 | -104.2595
21.006 | -104.2325
21.0098 | -104.206
21.0135 | -104.1797
21.0172 | -104.1538
21.0209 | -104.1309
21.0246 | -104.1097
21.0284 | -104.0905
21.0321 | -104.072
21.0358 | -104.0567
21.0395 | -104.0442
21.0432 | -104.0348
21.047 | -104.0273
21.0507 | -104.022
21.0544 | -104.0182
21.0581 | -104.0163
21.0618 | -104.0185
21.0656 | -104.0242
21.0693 | -104.0322
21.073 | -104.0429
21.0767 | -104.0564
21.0804 | -104.0728
21.0841 | -104.09
21.0879 | -104.1094
21.0916 | -104.1341
21.0953 | -104.1613
21.099 | -104.1874
21.1027 | -104.2175
21.1065 | -104.2523
21.1102 | -104.2882
21.1139 | -104.3243
21.1176 | -104.3609
21.1213 | -104.3983
21.1251 | -104.4385
21.1288 | -104.4796
21.1325 | -104.5227
21.1362 | -104.5658
21.1399 | -104.6066
21.1437 | -104.6496
21.1474 | -104.6938
21.1511 | -104.7357
21.1548 | -104.7755
21.1585 | -104.8172
21.1623 | -104.8602
21.166 | -104.9021
21.1697 | -104.9419
21.1734 | -104.9825
21.1771 | -105.0225
21.1808 | -105.0606
21.1846 | -105.095
21.1883 | -105.129
21.192 | -105.1589
21.1957 | -105.1842
21.1994 | -105.2109
21.2032 | -105.2357
21.2069 | -105.2602
21.2106 | -105.2845
21.2143 | -105.3074
21.218 | -105.3245
21.2218 | -105.333
21.2255 | -105.3433
21.2292 | -105.3522
21.2329 | -105.3559
21.2366 | -105.3611
21.2404 | -105.3635
21.2441 | -105.3641
21.2478 | -105.3637
21.2515 | -105.3619
21.2552 | -105.357
21.2589 | -105.3506
21.2627 | -105.343
21.2664 | -105.3319
21.2701 | -105.3203
21.2738 | -105.3061
21.2775 | -105.2886
21.2813 | -105.2706
21.285 | -105.2537
21.2887 | -105.2358
21.2924 | -105.2172
21.2961 | -105.1994
21.2999 | -105.1777
21.3036 | -105.1576
21.3073 | -105.1339
21.311 | -105.1115
21.3147 | -105.09
21.3185 | -105.0704
21.3222 | -105.0487
21.3259 | -105.0284
21.3296 | -105.0076
21.3333 | -104.9887
21.3371 | -104.9696
21.3408 | -104.9484
21.3445 | -104.9303
21.3482 | -104.9138
21.3519 | -104.8971
21.3556 | -104.8813
21.3594 | -104.8675
21.3631 | -104.8537
21.3668 | -104.8412
21.3705 | -104.8307
21.3742 | -104.8216
21.378 | -104.8122
21.3817 | -104.8043
21.3854 | -104.8004
21.3891 | -104.798
21.3928 | -104.7962
21.3966 | -104.796
21.4003 | -104.7996
21.404 | -104.8038
21.4077 | -104.8089
21.4114 | -104.8128
21.4152 | -104.8063
21.4189 | -104.7575
21.4226 | -104.6633
21.4263 | -104.5664
21.43 | -104.4526
21.4338 | -104.3398
21.4375 | -104.2409
21.4412 | -104.1502
21.4449 | -104.0663
21.4486 | -103.9808
21.4523 | -103.8898
21.4561 | -103.7918
21.4598 | -103.619
21.4635 | -103.5069
21.4672 | -103.389
21.4709 | -103.2652
21.4747 | -103.1375
21.4784 | -103.0068
21.4821 | -102.8721
21.4858 | -102.7349
21.4895 | -102.5967
21.4933 | -102.4574
21.497 | -102.3186
21.5007 | -102.1807
21.5044 | -102.0418
21.5081 | -101.9056
21.5119 | -101.773
21.5156 | -101.6433
21.5193 | -101.518
21.523 | -101.3988
21.5267 | -101.2865
21.5305 | -101.1791
21.5342 | -101.077
21.5379 | -100.9828
21.5416 | -100.8954
21.5453 | -100.8166
21.549 | -100.7456
21.5528 | -100.6836
21.5565 | -100.6298
21.5602 | -100.5854
21.5639 | -100.5496
21.5676 | -100.5234
21.5714 | -100.5069
21.5751 | -100.4998
21.5788 | -100.5021
21.5825 | -100.5149
21.5862 | -100.5366
21.59 | -100.5681
21.5937 | -100.6085
21.5974 | -100.6581
21.6011 | -100.7172
21.6048 | -100.7859
21.6086 | -100.8632
21.6123 | -100.9502
21.616 | -101.0438
21.6197 | -101.1463
21.6234 | -101.2557
21.6272 | -101.3728
21.6309 | -101.4964
21.6346 | -101.6267
21.6383 | -101.7629
21.642 | -101.904
21.6457 | -102.0504
21.6495 | -102.2006
21.6532 | -102.3536
21.6569 | -102.5092
21.6606 | -102.6677
21.6643 | -102.8271
21.6681 | -102.9876
21.6718 | -103.1472
21.6755 | -103.3057
21.6792 | -103.462
21.6829 | -103.6168
21.6867 | -103.7692
21.6904 | -103.9149
21.6941 | -104.0562
21.6978 | -104.1918
21.7015 | -104.3234
21.7053 | -104.4475
21.709 | -104.563
21.7127 | -104.6714
21.7164 | -104.7766
21.7201 | -104.883
21.7238 | -105.0196
21.7276 | -105.183
21.7313 | -105.3233
21.735 | -105.457
21.7387 | -105.5605
21.7424 | -105.6222
21.7462 | -105.6565
21.7499 | -105.6619
21.7536 | -105.6592
21.7573 | -105.6559
21.761 | -105.653
21.7648 | -105.6497
21.7685 | -105.643
21.7722 | -105.6353
21.7759 | -105.6278
21.7796 | -105.62
21.7834 | -105.6122
21.7871 | -105.6002
21.7908 | -105.5884
21.7945 | -105.5782
21.7982 | -105.5662
21.802 | -105.5544
21.8057 | -105.5432
21.8094 | -105.5322
21.8131 | -105.5206
21.8168 | -105.51
21.8205 | -105.4966
21.8243 | -105.4848
21.828 | -105.4706
21.8317 | -105.4546
21.8354 | -105.4391
21.8391 | -105.4223
21.8429 | -105.4062
21.8466 | -105.3891
21.8503 | -105.3749
21.854 | -105.3585
21.8577 | -105.3405
21.8615 | -105.3215
21.8652 | -105.3044
21.8689 | -105.2882
21.8726 | -105.2705
21.8763 | -105.2526
21.8801 | -105.2348
21.8838 | -105.2174
21.8875 | -105.2007
21.8912 | -105.1831
21.8949 | -105.1667
21.8987 | -105.1487
21.9024 | -105.132
21.9061 | -105.117
21.9098 | -105.0999
21.9135 | -105.0822
21.9172 | -105.0666
21.921 | -105.0501
21.9247 | -105.037
21.9284 | -105.0245
21.9321 | -105.0133
21.9358 | -105.0008
21.9396 | -104.9891
21.9433 | -104.9794
21.947 | -104.9695
21.9507 | -104.9613
21.9544 | -104.953
21.9582 | -104.9456
21.9619 | -104.9403
21.9656 | -104.9388
21.9693 | -104.9369
21.973 | -104.935
21.9768 | -104.9336
21.9805 | -104.9345
21.9842 | -104.9354
21.9879 | -104.9385
21.9916 | -104.9418
21.9954 | -104.9467
21.9991 | -104.9602
22.0028 | -104.9672
22.0065 | -104.9753
22.0102 | -104.9857
22.0139 | -104.9956
22.0177 | -105.0067
22.0214 | -105.021
22.0251 | -105.0357
22.0288 | -105.049
22.0325 | -105.0639
22.0363 | -105.0786
22.04 | -105.0936
22.0437 | -105.109
22.0474 | -105.1256
22.0511 | -105.1427
22.0549 | -105.1595
22.0586 | -105.1788
22.0623 | -105.2005
22.066 | -105.2195
22.0697 | -105.2387
22.0735 | -105.2576
22.0772 | -105.2777
22.0809 | -105.2987
22.0846 | -105.3181
22.0883 | -105.3391
22.0921 | -105.3594
22.0958 | -105.3746
22.0995 | -105.3939
22.1032 | -105.4132
22.1069 | -105.4284
22.1106 | -105.4455
22.1144 | -105.4628
22.1181 | -105.4789
22.1218 | -105.4932
22.1255 | -105.5093
22.1292 | -105.523
22.133 | -105.5353
22.1367 | -105.5468
22.1404 | -105.5588
22.1441 | -105.5724
22.1478 | -105.5856
22.1516 | -105.5977
22.1553 | -105.6093
22.159 | -105.6222
22.1627 | -105.6316
22.1664 | -105.6408
22.1702 | -105.649
22.1739 | -105.6576
22.1776 | -105.6659
22.1813 | -105.6722
22.185 | -105.6799
22.1887 | -105.6858
22.1925 | -105.6904
22.1962 | -105.696
22.1999 | -105.7009
22.2036 | -105.7027
22.2073 | -105.7042
22.2111 | -105.7044
22.2148 | -105.7051
22.2185 | -105.7034
22.2222 | -105.7014
22.2259 | -105.701
22.2297 | -105.698
22.2334 | -105.6958
22.2371 | -105.6892
22.2408 | -105.6806
22.2445 | -105.6709
22.2483 | -105.6647
22.252 | -105.6549
22.2557 | -105.6407
22.2594 | -105.6288
22.2631 | -105.6163
22.2669 | -105.6039
22.2706 | -105.5893
22.2743 | -105.5718
22.278 | -105.5503
22.2817 | -105.5313
22.2854 | -105.5113
22.2892 | -105.491
22.2929 | -105.4696
22.2966 | -105.4487
22.3003 | -105.4271
22.304 | -105.4039
22.3078 | -105.3784
22.3115 | -105.3539
22.3152 | -105.3307
22.3189 | -105.3091
22.3226 | -105.2866
22.3264 | -105.2631
22.3301 | -105.24
22.3338 | -105.2198
22.3375 | -105.1981
22.3412 | -105.1761
22.345 | -105.1572
22.3487 | -105.1392
22.3524 | -105.1243
22.3561 | -105.1101
22.3598 | -105.0944
22.3636 | -105.0784
22.3673 | -105.0614
22.371 | -105.0472
22.3747 | -105.0356
22.3784 | -105.0247
22.3821 | -105.0148
22.3859 | -105.0056
22.3896 | -104.9986
22.3933 | -104.9926
22.397 | -104.9861
22.4007 | -104.9817
22.4045 | -104.9784
22.4082 | -104.9745
22.4119 | -104.9728
22.4156 | -104.9685
22.4193 | -104.9666
22.4231 | -104.9675
22.4268 | -104.9663
22.4305 | -104.9646
22.4342 | -104.9648
22.4379 | -104.9651
22.4417 | -104.9651
22.4454 | -104.9655
22.4491 | -104.9672
22.4528 | -104.9699
22.4565 | -104.9699
22.4603 | -104.967
22.464 | -104.9642
22.4677 | -104.9628
22.4714 | -104.9612
22.4751 | -104.9589
22.4788 | -104.9545
22.4826 | -104.9513
22.4863 | -104.9507
22.49 | -104.9479
22.4937 | -104.9421
22.4974 | -104.9374
22.5012 | -104.932
22.5049 | -104.9287
22.5086 | -104.9268
22.5123 | -104.9228
22.516 | -104.9186
22.5198 | -104.9156
22.5235 | -104.9116
22.5272 | -104.9098
22.5309 | -104.9063
22.5346 | -104.9062
22.5384 | -104.9045
22.5421 | -104.9031
22.5458 | -104.9005
22.5495 | -104.8992
22.5532 | -104.8971
22.557 | -104.8977
22.5607 | -104.898
22.5644 | -104.899
22.5681 | -104.8978
22.5718 | -104.8998
22.5755 | -104.9017
22.5793 | -104.902
22.583 | -104.901
22.5867 | -104.8999
22.5904 | -104.8989
22.5941 | -104.9012
22.5979 | -104.9051
22.6016 | -104.9054
22.6053 | -104.9046
22.609 | -104.9045
22.6127 | -104.9033
22.6165 | -104.9019
22.6202 | -104.901
22.6239 | -104.9023
22.6276 | -104.9035
22.6313 | -104.9023
22.6351 | -104.898
22.6388 | -104.8919
22.6425 | -104.886
22.6462 | -104.8808
22.6499 | -104.8743
22.6536 | -104.8668
22.6574 | -104.86
22.6611 | -104.8547
22.6648 | -104.8475
22.6685 | -104.8402
22.6722 | -104.8303
22.676 | -104.8203
22.6797 | -104.8111
22.6834 | -104.8021
22.6871 | -104.7973
22.6908 | -104.7902
22.6946 | -104.7791
22.6983 | -104.7701
22.702 | -104.7621
22.7057 | -104.7536
22.7094 | -104.7434
22.7132 | -104.7345
22.7169 | -104.7292
22.7206 | -104.7241
22.7243 | -104.7178
22.728 | -104.7131
22.7318 | -104.7097
22.7355 | -104.7072
22.7392 | -104.7034
22.7429 | -104.6991
22.7466 | -104.6989
22.7503 | -104.6993
22.7541 | -104.6982
22.7578 | -104.6992
22.7615 | -104.7014
22.7652 | -104.7043
22.7689 | -104.7076
22.7727 | -104.711
22.7764 | -104.717
22.7801 | -104.7261
22.7838 | -104.7356
22.7875 | -104.7438
22.7913 | -104.7541
22.795 | -104.7643
22.7987 | -104.7765
22.8024 | -104.7886
22.8061 | -104.7996
22.8099 | -104.8116
22.8136 | -104.8256
22.8173 | -104.8391
22.821 | -104.8543
22.8247 | -104.8674
22.8285 | -104.8793
22.8322 | -104.8911
22.8359 | -104.9054
22.8396 | -104.9189
22.8433 | -104.934
22.847 | -104.9472
22.8508 | -104.9611
22.8545 | -104.9769
22.8582 | -104.992
22.8619 | -105.0071
22.8656 | -105.0216
22.8694 | -105.0352
22.8731 | -105.0486
22.8768 | -105.0629
22.8805 | -105.0774
22.8842 | -105.0919
22.888 | -105.1059
22.8917 | -105.1189
22.8954 | -105.131
22.8991 | -105.1448
22.9028 | -105.1571
22.9066 | -105.1695
22.9103 | -105.1841
22.914 | -105.198
22.9177 | -105.208
22.9214 | -105.2161
22.9252 | -105.2278
22.9289 | -105.2402
22.9326 | -105.2496
22.9363 | -105.2606
22.94 | -105.2698
22.9437 | -105.2752
22.9475 | -105.2809
22.9512 | -105.2881
22.9549 | -105.2941
22.9586 | -105.303
22.9623 | -105.3074
22.9661 | -105.3112
22.9698 | -105.3159
22.9735 | -105.3213
22.9772 | -105.3261
22.9809 | -105.3303
22.9847 | -105.3329
22.9884 | -105.335
22.9921 | -105.3363
22.9958 | -105.3379
22.9995 | -105.3375
23.0033 | -105.338
23.007 | -105.3389
23.0107 | -105.3368
23.0144 | -105.3356
23.0181 | -105.3365
23.0219 | -105.335
23.0256 | -105.3317
23.0293 | -105.3295
23.033 | -105.3281
23.0367 | -105.327
23.0404 | -105.3253
23.0442 | -105.3235
23.0479 | -105.3204
23.0516 | -105.3187
23.0553 | -105.3186
23.059 | -105.3178
23.0628 | -105.3157
23.0665 | -105.3166
23.0702 | -105.3245
23.0739 | -105.3261
23.0776 | -105.3278
23.0814 | -105.3333
23.0851 | -105.3393
23.0888 | -105.3431
23.0925 | -105.3464
23.0962 | -105.3518
23.1 | -105.357
23.1037 | -105.3644
23.1074 | -105.3711
23.1111 | -105.3768
23.1148 | -105.386
23.1185 | -105.3947
23.1223 | -105.4035
23.126 | -105.414
23.1297 | -105.4216
23.1334 | -105.4311
23.1371 | -105.4419
23.1409 | -105.4529
23.1446 | -105.4629
23.1483 | -105.4713
23.152 | -105.4802
23.1557 | -105.4901
23.1595 | -105.4992
23.1632 | -105.5064
23.1669 | -105.5144
23.1706 | -105.5221
23.1743 | -105.5308
23.1781 | -105.5375
23.1818 | -105.5417
23.1855 | -105.5451
23.1892 | -105.5482
23.1929 | -105.5508
23.1967 | -105.5536
23.2004 | -105.5542
23.2041 | -105.5529
23.2078 | -105.5521
23.2115 | -105.5502
23.2152 | -105.5465
23.219 | -105.5426
23.2227 | -105.5388
23.2264 | -105.5344
23.2301 | -105.5305
23.2338 | -105.5255
23.2376 | -105.5194
23.2413 | -105.5136
23.245 | -105.5087
23.2487 | -105.5051
23.2524 | -105.4977
23.2562 | -105.4915
23.2599 | -105.4864
23.2636 | -105.4804
23.2673 | -105.4752
23.271 | -105.4717
23.2748 | -105.4648
23.2785 | -105.4578
23.2822 | -105.4537
23.2859 | -105.4477
23.2896 | -105.4422
23.2934 | -105.4376
23.2971 | -105.4326
23.3008 | -105.4272
23.3045 | -105.4208
23.3082 | -105.4144
23.3119 | -105.4091
23.3157 | -105.4045
23.3194 | -105.4003
23.3231 | -105.3979
23.3268 | -105.3969
23.3305 | -105.3959
23.3343 | -105.395
23.338 | -105.3951
23.3417 | -105.3932
23.3454 | -105.3906
23.3491 | -105.3892
23.3529 | -105.3882
23.3566 | -105.3876
23.3603 | -105.3871
23.364 | -105.3854
23.3677 | -105.3854
23.3715 | -105.3861
23.3752 | -105.3895
23.3789 | -105.3914
23.3826 | -105.3912
23.3863 | -105.3927
23.3901 | -105.3955
23.3938 | -105.3965
23.3975 | -105.397
23.4012 | -105.3961
23.4049 | -105.3959
23.4086 | -105.3965
23.4124 | -105.3962
23.4161 | -105.3963
23.4198 | -105.3998
23.4235 | -105.4018
23.4272 | -105.4012
23.431 | -105.4019
23.4347 | -105.4014
23.4384 | -105.4006
23.4421 | -105.3996
23.4458 | -105.3983
23.4496 | -105.3976
23.4533 | -105.3964
23.457 | -105.3932
23.4607 | -105.3922
23.4644 | -105.3893
23.4682 | -105.3868
23.4719 | -105.3842
23.4756 | -105.3836
23.4793 | -105.3826
23.483 | -105.3816
23.4868 | -105.3804
23.4905 | -105.3803
23.4942 | -105.3779
23.4979 | -105.3752
23.5016 | -105.3719
23.5053 | -105.3698
23.5091 | -105.3676
23.5128 | -105.3675
23.5165 | -105.367
23.5202 | -105.3656
23.5239 | -105.3661
23.5277 | -105.3668
23.5314 | -105.3687
23.5351 | -105.371
23.5388 | -105.373
23.5425 | -105.3764
23.5463 | -105.3817
23.55 | -105.3863
23.5537 | -105.3905
23.5574 | -105.3945
23.5611 | -105.3991
23.5649 | -105.4065
23.5686 | -105.412
23.5723 | -105.4176
23.576 | -105.4245
23.5797 | -105.4308
23.5834 | -105.4397
23.5872 | -105.4488
23.5909 | -105.4586
23.5946 | -105.4677
23.5983 | -105.4757
23.602 | -105.486
23.6058 | -105.5007
23.6095 | -105.5115
23.6132 | -105.5222
23.6169 | -105.5333
23.6206 | -105.5467
23.6244 | -105.5629
23.6281 | -105.5774
23.6318 | -105.5911
23.6355 | -105.6047
23.6392 | -105.6173
23.643 | -105.6324
23.6467 | -105.6484
23.6504 | -105.6646
23.6541 | -105.6802
23.6578 | -105.6957
23.6616 | -105.7114
23.6653 | -105.7268
23.669 | -105.7398
23.6727 | -105.7533
23.6764 | -105.7696
23.6801 | -105.7853
23.6839 | -105.8009
23.6876 | -105.8169
23.6913 | -105.8321
23.695 | -105.8477
23.6987 | -105.863
23.7025 | -105.8757
23.7062 | -105.8879
23.7099 | -105.8998
23.7136 | -105.9139
23.7173 | -105.9292
23.7211 | -105.9417
23.7248 | -105.9546
23.7285 | -105.9652
23.7322 | -105.973
23.7359 | -105.9825
23.7397 | -105.9925
23.7434 | -106.0023
23.7471 | -106.0105
23.7508 | -106.0164
23.7545 | -106.019
23.7583 | -106.0261
23.762 | -106.034
23.7657 | -106.0391
23.7694 | -106.0428
23.7731 | -106.045
23.7768 | -106.0449
23.7806 | -106.0452
23.7843 | -106.0438
23.788 | -106.0433
23.7917 | -106.0429
23.7954 | -106.0408
23.7992 | -106.039
23.8029 | -106.0352
23.8066 | -106.0301
23.8103 | -106.0243
23.814 | -106.0188
23.8178 | -106.0154
23.8215 | -106.0117
23.8252 | -106.0069
23.8289 | -106.0013
23.8326 | -105.9945
23.8364 | -105.9877
23.8401 | -105.9787
23.8438 | -105.9707
23.8475 | -105.9621
23.8512 | -105.9545
23.855 | -105.9473
23.8587 | -105.9389
23.8624 | -105.9292
23.8661 | -105.9176
23.8698 | -105.9084
23.8735 | -105.8992
23.8773 | -105.8888
23.881 | -105.879
23.8847 | -105.8721
23.8884 | -105.8657
23.8921 | -105.8597
23.8959 | -105.8541
23.8996 | -105.8463
23.9033 | -105.8392
23.907 | -105.8349
23.9107 | -105.831
23.9145 | -105.8273
23.9182 | -105.8238
23.9219 | -105.8203
23.9256 | -105.8174
23.9293 | -105.8165
23.9331 | -105.8174
23.9368 | -105.8178
23.9405 | -105.8168
23.9442 | -105.8183
23.9479 | -105.8213
23.9517 | -105.8235
23.9554 | -105.8259
23.9591 | -105.8288
23.9628 | -105.8316
23.9665 | -105.8349
23.9702 | -105.8382
23.974 | -105.8433
23.9777 | -105.8518
23.9814 | -105.8585
23.9851 | -105.8645
23.9888 | -105.8724
23.9926 | -105.881
23.9963 | -105.8895
24 | -105.8986
24.0037 | -105.9067
24.0074 | -105.9172
24.0112 | -105.927
24.0149 | -105.9359
24.0186 | -105.9472
24.0223 | -105.9575
24.026 | -105.968
24.0298 | -105.9783
24.0335 | -105.9886
24.0372 | -105.9992
24.0409 | -106.0111
24.0446 | -106.023
24.0483 | -106.0345
24.0521 | -106.0451
24.0558 | -106.056
24.0595 | -106.0691
24.0632 | -106.0813
24.0669 | -106.0943
24.0707 | -106.1103
24.0744 | -106.1244
24.0781 | -106.1368
24.0818 | -106.1491
24.0855 | -106.1623
24.0893 | -106.1762
24.093 | -106.1907
24.0967 | -106.2013
24.1004 | -106.2126
24.1041 | -106.2261
24.1079 | -106.2407
24.1116 | -106.2539
24.1153 | -106.2674
24.119 | -106.2784
24.1227 | -106.289
24.1265 | -106.2986
24.1302 | -106.3081
24.1339 | -106.32
24.1376 | -106.3309
24.1413 | -106.3409
24.145 | -106.3461
24.1488 | -106.3545
24.1525 | -106.3614
24.1562 | -106.3675
24.1599 | -106.3717
24.1636 | -106.3734
24.1674 | -106.3765
24.1711 | -106.382
24.1748 | -106.3864
24.1785 | -106.3894
24.1822 | -106.3927
24.186 | -106.3947
24.1897 | -106.3955
24.1934 | -106.3966
24.1971 | -106.3959
24.2008 | -106.3923
24.2046 | -106.3894
24.2083 | -106.3885
24.212 | -106.3839
24.2157 | -106.3781
24.2194 | -106.373
24.2232 | -106.367
24.2269 | -106.3591
24.2306 | -106.3508
24.2343 | -106.3438
24.238 | -106.3363
24.2417 | -106.3297
24.2455 | -106.3223
24.2492 | -106.3135
24.2529 | -106.3052
24.2566 | -106.2995
24.2603 | -106.2913
24.2641 | -106.2839
24.2678 | -106.2776
24.2715 | -106.2714
24.2752 | -106.2649
24.2789 | -106.2585
24.2827 | -106.2526
24.2864 | -106.2479
24.2901 | -106.243
24.2938 | -106.2352
24.2975 | -106.224
24.3013 | -106.2149
24.305 | -106.2062
24.3087 | -106.1996
24.3124 | -106.1945
24.3161 | -106.1889
24.3199 | -106.1832
24.3236 | -106.1773
24.3273 | -106.172
24.331 | -106.1671
24.3347 | -106.163
24.3384 | -106.159
24.3422 | -106.1547
24.3459 | -106.1515
24.3496 | -106.147
24.3533 | -106.1427
24.357 | -106.1382
24.3608 | -106.133
24.3645 | -106.1269
24.3682 | -106.1211
24.3719 | -106.1163
24.3756 | -106.1113
24.3794 | -106.1033
24.3831 | -106.0952
24.3868 | -106.0875
24.3905 | -106.0801
24.3942 | -106.0713
24.398 | -106.0615
24.4017 | -106.0508
24.4054 | -106.0407
24.4091 | -106.031
24.4128 | -106.0207
24.4166 | -106.0114
24.4203 | -106.0002
24.424 | -105.9911
24.4277 | -105.9819
24.4314 | -105.9728
24.4351 | -105.9647
24.4389 | -105.9572
24.4426 | -105.9487
24.4463 | -105.9372
24.45 | -105.9261
24.4537 | -105.9162
24.4575 | -105.9075
24.4612 | -105.8984
24.4649 | -105.8914
24.4686 | -105.8828
24.4723 | -105.8722
24.4761 | -105.8616
24.4798 | -105.8526
24.4835 | -105.8442
24.4872 | -105.8359
24.4909 | -105.8265
24.4947 | -105.8179
24.4984 | -105.8102
24.5021 | -105.8035
24.5058 | -105.7978
24.5095 | -105.7933
24.5132 | -105.7886
24.517 | -105.7833
24.5207 | -105.7794
24.5244 | -105.7758
24.5281 | -105.7721
24.5318 | -105.7679
24.5356 | -105.7636
24.5393 | -105.759
24.543 | -105.7567
24.5467 | -105.7562
24.5504 | -105.7535
24.5542 | -105.7501
24.5579 | -105.7489
24.5616 | -105.7491
24.5653 | -105.7488
24.569 | -105.7459
24.5728 | -105.7411
24.5765 | -105.7361
24.5802 | -105.73
24.5839 | -105.7175
24.5876 | -105.6982
24.5914 | -105.658
24.5951 | -105.6028
24.5988 | -105.5392
24.6025 | -105.471
24.6062 | -105.4034
24.6099 | -105.3459
24.6137 | -105.286
24.6174 | -105.2227
24.6211 | -105.1584
24.6248 | -105.0912
24.6285 | -105.0217
24.6323 | -104.9514
24.636 | -104.881
24.6397 | -104.8087
24.6434 | -104.7344
24.6471 | -104.6548
24.6509 | -104.5714
24.6546 | -104.4838
24.6583 | -104.3936
24.662 | -104.3011
24.6657 | -104.2054
24.6695 | -104.1079
24.6732 | -104.0091
24.6769 | -103.9085
24.6806 | -103.7872
24.6843 | -103.6896
24.6881 | -103.5927
24.6918 | -103.4967
24.6955 | -103.4021
24.6992 | -103.3094
24.7029 | -103.2181
24.7066 | -103.1287
24.7104 | -103.0422
24.7141 | -102.9596
24.7178 | -102.8804
24.7215 | -102.8039
24.7252 | -102.7312
24.729 | -102.6639
24.7327 | -102.6009
24.7364 | -102.5426
24.7401 | -102.4889
24.7438 | -102.4402
24.7476 | -102.3974
24.7513 | -102.3602
24.755 | -102.3285
24.7587 | -102.3019
24.7624 | -102.2816
24.7662 | -102.2669
24.7699 | -102.2591
24.7736 | -102.2571
24.7773 | -102.2613
24.781 | -102.2714
24.7848 | -102.2867
24.7885 | -102.3086
24.7922 | -102.3366
24.7959 | -102.371
24.7996 | -102.4112
24.8033 | -102.4574
24.8071 | -102.5093
24.8108 | -102.5665
24.8145 | -102.6276
24.8182 | -102.6947
24.8219 | -102.7676
24.8257 | -102.8443
24.8294 | -102.9242
24.8331 | -103.0083
24.8368 | -103.0959
24.8405 | -103.1878
24.8443 | -103.283
24.848 | -103.381
24.8517 | -103.4811
24.8554 | -103.5822
24.8591 | -103.6849
24.8629 | -103.789
24.8666 | -103.8939
24.8703 | -103.9991
24.874 | -104.1034
24.8777 | -104.2067
24.8815 | -104.3105
24.8852 | -104.4129
24.8889 | -104.5143
24.8926 | -104.6133
24.8963 | -104.7094
24.9 | -104.8037
24.9038 | -104.8953
24.9075 | -104.982
24.9112 | -105.0656
24.9149 | -105.1459
24.9186 | -105.2239
24.9224 | -105.2992
24.9261 | -105.3711
24.9298 | -105.4371
24.9335 | -105.5042
24.9372 | -105.5725
24.941 | -105.6528
24.9447 | -105.7363
24.9484 | -105.8155
24.9521 | -105.8845
24.9558 | -105.941
24.9596 | -105.9758
24.9633 | -106.0057
24.967 | -106.031
24.9707 | -106.0495
24.9744 | -106.0639
24.9781 | -106.0753
24.9819 | -106.0813
24.9856 | -106.0836
24.9893 | -106.0871
24.993 | -106.0897
24.9967 | -106.0915
25.0005 | -106.0939
25.0042 | -106.0963
25.0079 | -106.0973
25.0116 | -106.0979
25.0153 | -106.1012
25.0191 | -106.1044
25.0228 | -106.1071
25.0265 | -106.1116
25.0302 | -106.1169
25.0339 | -106.1218
25.0377 | -106.1248
25.0414 | -106.1274
25.0451 | -106.1306
25.0488 | -106.1319
25.0525 | -106.1337
25.0563 | -106.1363
25.06 | -106.1424
25.0637 | -106.1463
25.0674 | -106.1502
25.0711 | -106.1532
25.0748 | -106.1571
25.0786 | -106.1619
25.0823 | -106.1653
25.086 | -106.1685
25.0897 | -106.1724
25.0934 | -106.1759
25.0972 | -106.1805
25.1009 | -106.1829
25.1046 | -106.184
25.1083 | -106.1855
25.112 | -106.1872
25.1158 | -106.1873
25.1195 | -106.1878
25.1232 | -106.1897
25.1269 | -106.1886
25.1306 | -106.187
25.1344 | -106.1859
25.1381 | -106.1865
25.1418 | -106.1854
25.1455 | -106.1844
25.1492 | -106.1838
25.153 | -106.1823
25.1567 | -106.1796
25.1604 | -106.1765
25.1641 | -106.1736
25.1678 | -106.1706
25.1715 | -106.1666
25.1753 | -106.1623
25.179 | -106.1599
25.1827 | -106.1574
25.1864 | -106.154
25.1901 | -106.1518
25.1939 | -106.1482
25.1976 | -106.1435
25.2013 | -106.1393
25.205 | -106.1353
25.2087 | -106.1323
25.2125 | -106.1281
25.2162 | -106.1267
25.2199 | -106.1237
25.2236 | -106.1217
25.2273 | -106.1204
25.2311 | -106.1186
25.2348 | -106.1173
25.2385 | -106.1156
25.2422 | -106.1144
25.2459 | -106.1117
25.2497 | -106.1087
25.2534 | -106.1081
25.2571 | -106.1081
25.2608 | -106.1073
25.2645 | -106.1066
25.2682 | -106.1045
25.272 | -106.1026
25.2757 | -106.1017
25.2794 | -106.1014
25.2831 | -106.0994
25.2868 | -106.1001
25.2906 | -106.1006
25.2943 | -106.1006
25.298 | -106.101
25.3017 | -106.1005
25.3054 | -106.101
25.3092 | -106.1014
25.3129 | -106.1036
25.3166 | -106.1066
25.3203 | -106.1059
25.324 | -106.1063
25.3278 | -106.1075
25.3315 | -106.1065
25.3352 | -106.1057
25.3389 | -106.104
25.3426 | -106.1025
25.3464 | -106.1011
25.3501 | -106.1012
25.3538 | -106.1016
25.3575 | -106.1028
25.3612 | -106.1044
25.3649 | -106.1049
25.3687 | -106.1061
25.3724 | -106.1061
25.3761 | -106.1077
25.3798 | -106.1086
25.3835 | -106.1095
25.3873 | -106.1104
25.391 | -106.1103
25.3947 | -106.1114
25.3984 | -106.1136
25.4021 | -106.1155
25.4059 | -106.1195
25.4096 | -106.1235
25.4133 | -106.128
25.417 | -106.1289
25.4207 | -106.1296
25.4245 | -106.1308
25.4282 | -106.1338
25.4319 | -106.1376
25.4356 | -106.1391
25.4393 | -106.1402
25.443 | -106.1434
25.4468 | -106.1465
25.4505 | -106.1502
25.4542 | -106.1545
25.4579 | -106.1593
25.4616 | -106.166
25.4654 | -106.1715
25.4691 | -106.1777
25.4728 | -106.1841
25.4765 | -106.1891
25.4802 | -106.1921
25.484 | -106.1969
25.4877 | -106.2011
25.4914 | -106.2061
25.4951 | -106.2108
25.4988 | -106.216
25.5026 | -106.22
25.5063 | -106.2242
25.51 | -106.2284
25.5137 | -106.2329
25.5174 | -106.2359
25.5212 | -106.2387
25.5249 | -106.2426
25.5286 | -106.2462
25.5323 | -106.2493
25.536 | -106.2493
25.5397 | -106.2506
25.5435 | -106.2524
25.5472 | -106.2536
25.5509 | -106.2558
25.5546 | -106.257
25.5583 | -106.2564
25.5621 | -106.2545
25.5658 | -106.2538
25.5695 | -106.2544
25.5732 | -106.2548
25.5769 | -106.2535
25.5807 | -106.2515
25.5844 | -106.2497
25.5881 | -106.245
25.5918 | -106.2397
25.5955 | -106.2346
25.5993 | -106.2308
25.603 | -106.227
25.6067 | -106.223
25.6104 | -106.2194
25.6141 | -106.2143
25.6179 | -106.2084
25.6216 | -106.2025
25.6253 | -106.1966
25.629 | -106.1923
25.6327 | -106.1884
25.6364 | -106.1853
25.6402 | -106.1803
25.6439 | -106.1744
25.6476 | -106.168
25.6513 | -106.1621
25.655 | -106.1566
25.6588 | -106.1513
25.6625 | -106.1466
25.6662 | -106.1393
25.6699 | -106.1329
25.6736 | -106.1261
25.6774 | -106.1187
25.6811 | -106.1121
25.6848 | -106.1058
25.6885 | -106.1016
25.6922 | -106.099
25.696 | -106.0942
25.6997 | -106.0902
25.7034 | -106.0862
25.7071 | -106.0809
25.7108 | -106.0752
25.7146 | -106.0711
25.7183 | -106.0652
25.722 | -106.0604
25.7257 | -106.0573
25.7294 | -106.0536
25.7331 | -106.0507
25.7369 | -106.0472
25.7406 | -106.0427
25.7443 | -106.0386
25.748 | -106.0355
25.7517 | -106.0349
25.7555 | -106.0328
25.7592 | -106.031
25.7629 | -106.0283
25.7666 | -106.0265
25.7703 | -106.0243
25.7741 | -106.0211
25.7778 | -106.0186
25.7815 | -106.0176
25.7852 | -106.0179
25.7889 | -106.0186
25.7927 | -106.0195
25.7964 | -106.0187
25.8001 | -106.0195
25.8038 | -106.0227
25.8075 | -106.0239
25.8113 | -106.0253
25.815 | -106.027
25.8187 | -106.029
25.8224 | -106.032
25.8261 | -106.0331
25.8298 | -106.0364
25.8336 | -106.0399
25.8373 | -106.0431
25.841 | -106.0475
25.8447 | -106.0528
25.8484 | -106.0567
25.8522 | -106.0608
25.8559 | -106.0626
25.8596 | -106.0656
25.8633 | -106.0691
25.867 | -106.0725
25.8708 | -106.0759
25.8745 | -106.0781
25.8782 | -106.0796
25.8819 | -106.0831
25.8856 | -106.0864
25.8894 | -106.0893
25.8931 | -106.0915
25.8968 | -106.0935
25.9005 | -106.0958
25.9042 | -106.0973
25.9079 | -106.0979
25.9117 | -106.0991
25.9154 | -106.0988
25.9191 | -106.0993
25.9228 | -106.0997
25.9265 | -106.102
25.9303 | -106.1023
25.934 | -106.1016
25.9377 | -106.1004
25.9414 | -106.1003
25.9451 | -106.0987
25.9489 | -106.097
25.9526 | -106.098
25.9563 | -106.0996
25.96 | -106.1006
25.9637 | -106.0989
25.9675 | -106.0962
25.9712 | -106.0944
25.9749 | -106.093
25.9786 | -106.0923
25.9823 | -106.0912
25.9861 | -106.091
25.9898 | -106.091
25.9935 | -106.0891
25.9972 | -106.0859
26.0009 | -106.0833
26.0046 | -106.083
26.0084 | -106.0828
26.0121 | -106.0807
26.0158 | -106.0772
26.0195 | -106.0753
26.0232 | -106.0736
26.027 | -106.0731
26.0307 | -106.0726
26.0344 | -106.0704
26.0381 | -106.0683
26.0418 | -106.0672
26.0456 | -106.0677
26.0493 | -106.0675
26.053 | -106.0648
26.0567 | -106.0627
26.0604 | -106.0616
26.0642 | -106.0608
26.0679 | -106.0586
26.0716 | -106.0565
26.0753 | -106.0547
26.079 | -106.0519
26.0828 | -106.0505
26.0865 | -106.0499
26.0902 | -106.0486
26.0939 | -106.0484
26.0976 | -106.0456
26.1013 | -106.0439
26.1051 | -106.0441
26.1088 | -106.044
26.1125 | -106.0419
26.1162 | -106.0399
26.1199 | -106.0377
26.1237 | -106.0344
26.1274 | -106.0321
26.1311 | -106.0313
26.1348 | -106.0297
26.1385 | -106.0289
26.1423 | -106.0277
26.146 | -106.0263
26.1497 | -106.0251
26.1534 | -106.024
26.1571 | -106.0222
26.1609 | -106.0205
26.1646 | -106.0209
26.1683 | -106.0215
26.172 | -106.0208
26.1757 | -106.0205
26.1795 | -106.0215
26.1832 | -106.0238
26.1869 | -106.025
26.1906 | -106.0262
26.1943 | -106.0272
26.198 | -106.0264
26.2018 | -106.0276
26.2055 | -106.0277
26.2092 | -106.029
26.2129 | -106.032
26.2166 | -106.0349
26.2204 | -106.0394
26.2241 | -106.0453
26.2278 | -106.05
26.2315 | -106.0536
26.2352 | -106.0585
26.239 | -106.0627
26.2427 | -106.0686
26.2464 | -106.074
26.2501 | -106.0776
26.2538 | -106.0814
26.2576 | -106.0846
26.2613 | -106.0891
26.265 | -106.0927
26.2687 | -106.0957
26.2724 | -106.1
26.2762 | -106.105
26.2799 | -106.1094
26.2836 | -106.1144
26.2873 | -106.1174
26.291 | -106.119
26.2947 | -106.1234
26.2985 | -106.127
26.3022 | -106.1318
26.3059 | -106.1366
26.3096 | -106.1407
26.3133 | -106.1442
26.3171 | -106.15
26.3208 | -106.1548
26.3245 | -106.1591
26.3282 | -106.1627
26.3319 | -106.1671
26.3357 | -106.1692
26.3394 | -106.1722
26.3431 | -106.1759
26.3468 | -106.177
26.3505 | -106.1776
26.3543 | -106.1805
26.358 | -106.1824
26.3617 | -106.1845
26.3654 | -106.1874
26.3691 | -106.1888
26.3728 | -106.1891
26.3766 | -106.1902
26.3803 | -106.1918
26.384 | -106.1937
26.3877 | -106.196
26.3914 | -106.1975
26.3952 | -106.1984
26.3989 | -106.1991
26.4026 | -106.2012
26.4063 | -106.2038
26.41 | -106.2057
26.4138 | -106.2081
26.4175 | -106.2091
26.4212 | -106.2095
26.4249 | -106.2109
26.4286 | -106.2134
26.4324 | -106.2171
26.4361 | -106.2183
26.4398 | -106.2206
26.4435 | -106.2245
26.4472 | -106.2269
26.451 | -106.2298
26.4547 | -106.2304
26.4584 | -106.2299
26.4621 | -106.2324
26.4658 | -106.2344
26.4695 | -106.2373
26.4733 | -106.2397
26.477 | -106.2425
26.4807 | -106.2446
26.4844 | -106.2496
26.4881 | -106.2543
26.4919 | -106.2588
26.4956 | -106.2644
26.4993 | -106.2681
26.503 | -106.2693
26.5067 | -106.2709
26.5105 | -106.2724
26.5142 | -106.2747
26.5179 | -106.2781
26.5216 | -106.282
26.5253 | -106.2846
26.5291 | -106.2856
26.5328 | -106.286
26.5365 | -106.2864
26.5402 | -106.2863
26.5439 | -106.287
26.5477 | -106.2869
26.5514 | -106.2861
26.5551 | -106.2831
26.5588 | -106.2823
26.5625 | -106.2802
26.5662 | -106.2806
26.57 | -106.2809
26.5737 | -106.2778
26.5774 | -106.2755
26.5811 | -106.2705
26.5848 | -106.2644
26.5886 | -106.2592
26.5923 | -106.254
26.596 | -106.2482
26.5997 | -106.2404
26.6034 | -106.2331
26.6072 | -106.2258
26.6109 | -106.219
26.6146 | -106.2119
26.6183 | -106.205
26.622 | -106.1981
26.6258 | -106.1918
26.6295 | -106.1859
26.6332 | -106.1794
26.6369 | -106.173
26.6406 | -106.166
26.6444 | -106.1576
26.6481 | -106.1495
26.6518 | -106.1418
26.6555 | -106.1342
26.6592 | -106.1288
26.6629 | -106.1231
26.6667 | -106.1164
26.6704 | -106.1089
26.6741 | -106.1019
26.6778 | -106.0953
26.6815 | -106.0902
26.6853 | -106.0867
26.689 | -106.0834
26.6927 | -106.0778
26.6964 | -106.0729
26.7001 | -106.0693
26.7039 | -106.0677
26.7076 | -106.0659
26.7113 | -106.0671
26.715 | -106.0678
26.7187 | -106.0672
26.7225 | -106.0663
26.7262 | -106.0655
26.7299 | -106.063
26.7336 | -106.0606
26.7373 | -106.0574
26.7411 | -106.0576
26.7448 | -106.0591
26.7485 | -106.0593
26.7522 | -106.0597
26.7559 | -106.0613
26.7596 | -106.063
26.7634 | -106.0642
26.7671 | -106.0654
26.7708 | -106.0668
26.7745 | -106.0683
26.7782 | -106.0687
26.782 | -106.0708
26.7857 | -106.0727
26.7894 | -106.0732
26.7931 | -106.0744
26.7968 | -106.0771
26.8006 | -106.0807
26.8043 | -106.0814
26.808 | -106.0806
26.8117 | -106.0827
26.8154 | -106.0835
26.8192 | -106.0834
26.8229 | -106.0838
26.8266 | -106.0758
26.8303 | -106.0727
26.834 | -106.0712
26.8377 | -106.0692
26.8415 | -106.0636
26.8452 | -106.0593
26.8489 | -106.0568
26.8526 | -106.0551
26.8563 | -106.0521
26.8601 | -106.0487
26.8638 | -106.0435
26.8675 | -106.0386
26.8712 | -106.0331
26.8749 | -106.0274
26.8787 | -106.0242
26.8824 | -106.0215
26.8861 | -106.0185
26.8898 | -106.0139
26.8935 | -106.0096
26.8973 | -106.0056
26.901 | -106.0004
26.9047 | -105.9961
26.9084 | -105.9934
26.9121 | -105.9899
26.9159 | -105.9862
26.9196 | -105.9823
26.9233 | -105.978
26.927 | -105.9749
26.9307 | -105.9719
26.9344 | -105.9701
26.9382 | -105.9674
26.9419 | -105.9671
26.9456 | -105.9657
26.9493 | -105.9645
26.953 | -105.9644
26.9568 | -105.9633
26.9605 | -105.9644
26.9642 | -105.9651
26.9679 | -105.9653
26.9716 | -105.9647
26.9754 | -105.966
26.9791 | -105.9694
26.9828 | -105.9723
26.9865 | -105.9752
26.9902 | -105.9786
26.994 | -105.9818
26.9977 | -105.9865
27.0014 | -105.9905
27.0051 | -105.9929
27.0088 | -105.9957
27.0126 | -105.9997
27.0163 | -106.0054
27.02 | -106.0122
27.0237 | -106.0183
27.0274 | -106.0244
27.0311 | -106.0315
27.0349 | -106.0384
27.0386 | -106.0446
27.0423 | -106.0501
27.046 | -106.0565
27.0497 | -106.0633
27.0535 | -106.0695
27.0572 | -106.0753
27.0609 | -106.0804
27.0646 | -106.0853
27.0683 | -106.0918
27.0721 | -106.1
27.0758 | -106.1077
27.0795 | -106.116
27.0832 | -106.1218
27.0869 | -106.1275
27.0907 | -106.1338
27.0944 | -106.1396
27.0981 | -106.1458
27.1018 | -106.1541
27.1055 | -106.1627
27.1093 | -106.1709
27.113 | -106.178
27.1167 | -106.1844
27.1204 | -106.1877
27.1241 | -106.1917
27.1278 | -106.1956
27.1316 | -106.1977
27.1353 | -106.1992
27.139 | -106.2013
27.1427 | -106.201
27.1464 | -106.2003
27.1502 | -106.2018
27.1539 | -106.2019
27.1576 | -106.203
27.1613 | -106.2026
27.165 | -106.2025
27.1688 | -106.2014
27.1725 | -106.1991
27.1762 | -106.1971
27.1799 | -106.1941
27.1836 | -106.1918
27.1874 | -106.1903
27.1911 | -106.187
27.1948 | -106.184
27.1985 | -106.1815
27.2022 | -106.1774
27.206 | -106.1748
27.2097 | -106.171
27.2134 | -106.1673
27.2171 | -106.1646
27.2208 | -106.1617
27.2245 | -106.159
27.2283 | -106.1533
27.232 | -106.1468
27.2357 | -106.1402
27.2394 | -106.1335
27.2431 | -106.1286
27.2469 | -106.1246
27.2506 | -106.1198
27.2543 | -106.114
27.258 | -106.109
27.2617 | -106.1029
27.2655 | -106.0973
27.2692 | -106.0916
27.2729 | -106.0884
27.2766 | -106.0846
27.2803 | -106.0795
27.2841 | -106.0769
27.2878 | -106.0749
27.2915 | -106.0724
27.2952 | -106.0698
27.2989 | -106.067
27.3026 | -106.0659
27.3064 | -106.0642
27.3101 | -106.0625
27.3138 | -106.0623
27.3175 | -106.0625
27.3212 | -106.0623
27.325 | -106.0608
27.3287 | -106.0605
27.3324 | -106.0589
27.3361 | -106.0579
27.3398 | -106.0591
27.3436 | -106.0596
27.3473 | -106.0599
27.351 | -106.0619
27.3547 | -106.0644
27.3584 | -106.0649
27.3622 | -106.0678
27.3659 | -106.0692
27.3696 | -106.0709
27.3733 | -106.0724
27.377 | -106.0742
27.3808 | -106.0771
27.3845 | -106.0778
27.3882 | -106.0797
27.3919 | -106.0831
27.3956 | -106.0843
27.3993 | -106.0865
27.4031 | -106.0886
27.4068 | -106.0901
27.4105 | -106.0903
27.4142 | -106.0905
27.4179 | -106.0903
27.4217 | -106.0904
27.4254 | -106.0925
27.4291 | -106.0953
27.4328 | -106.0971
27.4365 | -106.0981
27.4403 | -106.1003
27.444 | -106.1014
27.4477 | -106.1012
27.4514 | -106.1011
27.4551 | -106.0993
27.4589 | -106.0981
27.4626 | -106.0982
27.4663 | -106.0949
27.47 | -106.0906
27.4737 | -106.0872
27.4775 | -106.085
27.4812 | -106.0824
27.4849 | -106.0806
27.4886 | -106.0779
27.4923 | -106.0737
27.496 | -106.0692
27.4998 | -106.0661
27.5035 | -106.0645
27.5072 | -106.0601
27.5109 | -106.0566
27.5146 | -106.0544
27.5184 | -106.0515
27.5221 | -106.0475
27.5258 | -106.0438
27.5295 | -106.0408
27.5332 | -106.0353
27.537 | -106.0273
27.5407 | -106.022
27.5444 | -106.0154
27.5481 | -106.0087
27.5518 | -106.0018
27.5556 | -105.9957
27.5593 | -105.9903
27.563 | -105.985
27.5667 | -105.9793
27.5704 | -105.9743
27.5742 | -105.9683
27.5779 | -105.9621
27.5816 | -105.9563
27.5853 | -105.9501
27.589 | -105.9458
27.5927 | -105.941
27.5965 | -105.9382
27.6002 | -105.9331
27.6039 | -105.9268
27.6076 | -105.9205
27.6113 | -105.915
27.6151 | -105.9104
27.6188 | -105.9033
27.6225 | -105.8961
27.6262 | -105.89
27.6299 | -105.8844
27.6337 | -105.8775
27.6374 | -105.8713
27.6411 | -105.8658
27.6448 | -105.8599
27.6485 | -105.8547
27.6523 | -105.8508
27.656 | -105.8474
27.6597 | -105.8429
27.6634 | -105.8386
27.6671 | -105.8356
27.6709 | -105.8318
27.6746 | -105.8293
27.6783 | -105.828
27.682 | -105.8266
27.6857 | -105.8255
27.6894 | -105.8246
27.6932 | -105.825
27.6969 | -105.8249
27.7006 | -105.823
27.7043 | -105.8217
27.708 | -105.8212
27.7118 | -105.8211
27.7155 | -105.8204
27.7192 | -105.8219
27.7229 | -105.8238
27.7266 | -105.8254
27.7304 | -105.8271
27.7341 | -105.8279
27.7378 | -105.829
27.7415 | -105.8304
27.7452 | -105.8315
27.749 | -105.8329
27.7527 | -105.8321
27.7564 | -105.8318
27.7601 | -105.8321
27.7638 | -105.8334
27.7675 | -105.8351
27.7713 | -105.8358
27.775 | -105.8387
27.7787 | -105.8422
27.7824 | -105.8454
27.7861 | -105.8489
27.7899 | -105.8498
27.7936 | -105.8513
27.7973 | -105.8539
27.801 | -105.858
27.8047 | -105.8608
27.8085 | -105.8628
27.8122 | -105.8654
27.8159 | -105.8672
27.8196 | -105.8687
27.8233 | -105.8704
27.8271 | -105.8727
27.8308 | -105.8734
27.8345 | -105.8751
27.8382 | -105.8787
27.8419 | -105.881
27.8457 | -105.8837
27.8494 | -105.8871
27.8531 | -105.8893
27.8568 | -105.8924
27.8605 | -105.8969
27.8642 | -105.9002
27.868 | -105.9015
27.8717 | -105.9034
27.8754 | -105.905
27.8791 | -105.9076
27.8828 | -105.9104
27.8866 | -105.914
27.8903 | -105.918
27.894 | -105.9202
27.8977 | -105.9229
27.9014 | -105.9269
27.9052 | -105.929
27.9089 | -105.9302
27.9126 | -105.9312
27.9163 | -105.9327
27.92 | -105.9337
27.9238 | -105.9343
27.9275 | -105.9363
27.9312 | -105.9382
27.9349 | -105.9371
27.9386 | -105.9364
27.9424 | -105.9361
27.9461 | -105.9362
27.9498 | -105.9367
27.9535 | -105.9362
27.9572 | -105.9339
27.9609 | -105.9306
27.9647 | -105.9263
27.9684 | -105.9234
27.9721 | -105.9219
27.9758 | -105.9195
27.9795 | -105.9172
27.9833 | -105.9138
27.987 | -105.9109
27.9907 | -105.9066
27.9944 | -105.9031
27.9981 | -105.8974
28.0019 | -105.8907