## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 212 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 210
Hood | 02-92B/P | 144
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 986 | hPa |  | PRC PRO 01
Humidity | 34.3 | %/Hr
Internal temperature of calibration bench | 23.63 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 995 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.6 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.91 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.03 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3143 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 30/03/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 06/04/2018
Engineer | Anthony HUE | Approval |  | 23/04/2018
CEO | Robert GLEMOT | Approval |  | 23/04/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5136 | 57.5387
0.012209 | -36.0862 | 38.3686
0.018313 | -35.0154 | 28.2902
0.024417 | -34.5511 | 22.0212
0.030521 | -34.4448 | 18.2543
0.036626 | -34.2999 | 15.1982
0.04273 | -34.2025 | 12.8241
0.048834 | -34.1554 | 11.3009
0.054938 | -34.0887 | 10.1718
0.061043 | -34.0275 | 9.1015
0.067147 | -34.0165 | 8.281
0.073251 | -34.0168 | 7.5174
0.079355 | -34.0355 | 6.8492
0.08546 | -34.051 | 6.2445
0.091564 | -34.047 | 5.7953
0.097668 | -34.0282 | 5.5492
0.10377 | -34.0054 | 5.305
0.10988 | -33.9879 | 5.0281
0.11598 | -33.9788 | 4.7342
0.12209 | -33.9809 | 4.4499
0.12819 | -33.9913 | 4.166
0.13429 | -34.0025 | 3.8767
0.1404 | -34.006 | 3.5948
0.1465 | -34.008 | 3.3344
0.15261 | -34.005 | 3.113
0.15871 | -33.998 | 2.952
0.16482 | -33.9894 | 2.8526
0.17092 | -33.9869 | 2.8178
0.17702 | -33.9845 | 2.7484
0.18313 | -33.9833 | 2.672
0.18923 | -33.9825 | 2.5628
0.19534 | -33.9807 | 2.4338
0.20144 | -33.9789 | 2.2874
0.20754 | -33.9755 | 2.1624
0.21365 | -33.9742 | 2.0445
0.21975 | -33.9736 | 1.9446
0.22586 | -33.9746 | 1.8563
0.23196 | -33.9771 | 1.7844
0.23807 | -33.9814 | 1.7235
0.24417 | -33.9857 | 1.6677
0.25027 | -33.9899 | 1.6072
0.25638 | -33.9923 | 1.5471
0.26248 | -33.9942 | 1.4814
0.26859 | -33.995 | 1.4169
0.27469 | -33.9953 | 1.3452
0.2808 | -33.9963 | 1.2689
0.2869 | -33.997 | 1.1961
0.293 | -33.9979 | 1.127
0.29911 | -33.9987 | 1.0649
0.30521 | -33.9997 | 1.0036
0.31132 | -34.0003 | 0.9638
0.31742 | -34.0015 | 0.91958
0.32353 | -34.0017 | 0.88385
0.32963 | -34.0025 | 0.85447
0.33573 | -34.0036 | 0.83463
0.34184 | -34.0049 | 0.81367
0.34794 | -34.0072 | 0.79096
0.35405 | -34.0095 | 0.76765
0.36015 | -34.0121 | 0.74327
0.36626 | -34.0143 | 0.71655
0.37236 | -34.0166 | 0.6803
0.37846 | -34.0189 | 0.6359
0.38457 | -34.0199 | 0.57743
0.39067 | -34.0211 | 0.53006
0.39678 | -34.0218 | 0.47701
0.40288 | -34.0223 | 0.42173
0.40899 | -34.0222 | 0.36854
0.41509 | -34.0215 | 0.31644
0.42119 | -34.0213 | 0.27249
0.4273 | -34.0212 | 0.22526
0.4334 | -34.0211 | 0.17822
0.43951 | -34.0214 | 0.13424
0.44561 | -34.0222 | 0.094124
0.45172 | -34.0239 | 0.062605
0.45782 | -34.0247 | 0.025298
0.46392 | -34.0259 | -0.011092
0.47003 | -34.027 | -0.04336
0.47613 | -34.0279 | -0.077798
0.48224 | -34.0284 | -0.11294
0.48834 | -34.0275 | -0.14847
0.49445 | -34.0272 | -0.18437
0.50055 | -34.0267 | -0.21683
0.50665 | -34.0259 | -0.25178
0.51276 | -34.0248 | -0.28657
0.51886 | -34.0238 | -0.32137
0.52497 | -34.0225 | -0.346
0.53107 | -34.0217 | -0.37613
0.53717 | -34.021 | -0.4049
0.54328 | -34.0201 | -0.42798
0.54938 | -34.0199 | -0.4507
0.55549 | -34.0199 | -0.47183
0.56159 | -34.0209 | -0.48314
0.5677 | -34.0212 | -0.4977
0.5738 | -34.0222 | -0.51395
0.5799 | -34.023 | -0.52887
0.58601 | -34.024 | -0.53863
0.59211 | -34.0256 | -0.5443
0.59822 | -34.0269 | -0.5534
0.60432 | -34.0284 | -0.56518
0.61043 | -34.0296 | -0.57261
0.61653 | -34.0308 | -0.58251
0.62263 | -34.0323 | -0.59134
0.62874 | -34.0329 | -0.60155
0.63484 | -34.0337 | -0.61033
0.64095 | -34.0341 | -0.62014
0.64705 | -34.0344 | -0.62945
0.65316 | -34.0345 | -0.64163
0.65926 | -34.0346 | -0.65736
0.66536 | -34.0334 | -0.67488
0.67147 | -34.0329 | -0.68792
0.67757 | -34.0325 | -0.70214
0.68368 | -34.0316 | -0.72128
0.68978 | -34.0308 | -0.73854
0.69589 | -34.0298 | -0.75775
0.70199 | -34.0287 | -0.78099
0.70809 | -34.0277 | -0.80039
0.7142 | -34.0269 | -0.82123
0.7203 | -34.0257 | -0.8407
0.72641 | -34.025 | -0.86369
0.73251 | -34.0243 | -0.88401
0.73862 | -34.0245 | -0.9063
0.74472 | -34.0241 | -0.92667
0.75082 | -34.0238 | -0.94923
0.75693 | -34.0241 | -0.97148
0.76303 | -34.0246 | -0.99399
0.76914 | -34.0265 | -1.015
0.77524 | -34.0275 | -1.035
0.78135 | -34.0288 | -1.0555
0.78745 | -34.0302 | -1.0759
0.79355 | -34.0315 | -1.0961
0.79966 | -34.0333 | -1.1159
0.80576 | -34.0357 | -1.1362
0.81187 | -34.0374 | -1.1584
0.81797 | -34.0394 | -1.1803
0.82408 | -34.0412 | -1.2013
0.83018 | -34.043 | -1.2225
0.83628 | -34.045 | -1.2414
0.84239 | -34.0462 | -1.2626
0.84849 | -34.0477 | -1.285
0.8546 | -34.0493 | -1.3054
0.8607 | -34.0506 | -1.3268
0.86681 | -34.0517 | -1.3462
0.87291 | -34.0527 | -1.3656
0.87901 | -34.0532 | -1.3842
0.88512 | -34.0539 | -1.4029
0.89122 | -34.0547 | -1.4217
0.89733 | -34.0551 | -1.4377
0.90343 | -34.0555 | -1.4558
0.90953 | -34.0557 | -1.4705
0.91564 | -34.0557 | -1.4867
0.92174 | -34.0559 | -1.502
0.92785 | -34.0559 | -1.5169
0.93395 | -34.056 | -1.5311
0.94006 | -34.0561 | -1.5442
0.94616 | -34.0558 | -1.5562
0.95226 | -34.0559 | -1.5691
0.95837 | -34.0559 | -1.5832
0.96447 | -34.0561 | -1.5962
0.97058 | -34.0559 | -1.6098
0.97668 | -34.0556 | -1.6234
0.98279 | -34.0553 | -1.6364
0.98889 | -34.0551 | -1.6471
0.99499 | -34.055 | -1.6607
1.0011 | -34.0548 | -1.6715
1.0072 | -34.0544 | -1.6832
1.0133 | -34.054 | -1.694
1.0194 | -34.0535 | -1.7071
1.0255 | -34.0531 | -1.721
1.0316 | -34.0525 | -1.733
1.0377 | -34.052 | -1.7468
1.0438 | -34.0514 | -1.7596
1.0499 | -34.0505 | -1.7751
1.056 | -34.0498 | -1.7896
1.0621 | -34.0491 | -1.804
1.0682 | -34.0485 | -1.8165
1.0743 | -34.0478 | -1.8322
1.0805 | -34.0473 | -1.8467
1.0866 | -34.0462 | -1.862
1.0927 | -34.0453 | -1.8766
1.0988 | -34.0446 | -1.8917
1.1049 | -34.0437 | -1.906
1.111 | -34.0428 | -1.9216
1.1171 | -34.0421 | -1.9363
1.1232 | -34.0415 | -1.9516
1.1293 | -34.0411 | -1.9657
1.1354 | -34.0407 | -1.9809
1.1415 | -34.0403 | -1.9966
1.1476 | -34.0396 | -2.0123
1.1537 | -34.0391 | -2.0271
1.1598 | -34.0391 | -2.042
1.1659 | -34.0389 | -2.0554
1.172 | -34.0386 | -2.0694
1.1781 | -34.0383 | -2.0835
1.1842 | -34.0381 | -2.0978
1.1903 | -34.0385 | -2.1088
1.1964 | -34.0383 | -2.1223
1.2025 | -34.0384 | -2.1364
1.2086 | -34.0388 | -2.1488
1.2147 | -34.0388 | -2.1614
1.2209 | -34.0391 | -2.1726
1.227 | -34.0399 | -2.1833
1.2331 | -34.0403 | -2.1954
1.2392 | -34.0407 | -2.2073
1.2453 | -34.0411 | -2.2184
1.2514 | -34.0415 | -2.2282
1.2575 | -34.0421 | -2.2392
1.2636 | -34.043 | -2.2482
1.2697 | -34.0437 | -2.2574
1.2758 | -34.0444 | -2.2686
1.2819 | -34.0451 | -2.2786
1.288 | -34.0458 | -2.2894
1.2941 | -34.0467 | -2.2978
1.3002 | -34.0474 | -2.3058
1.3063 | -34.0479 | -2.3148
1.3124 | -34.0486 | -2.3238
1.3185 | -34.0494 | -2.3329
1.3246 | -34.0501 | -2.342
1.3307 | -34.051 | -2.3506
1.3368 | -34.0518 | -2.3592
1.3429 | -34.0525 | -2.3682
1.349 | -34.053 | -2.3775
1.3551 | -34.0537 | -2.3869
1.3613 | -34.0543 | -2.3951
1.3674 | -34.055 | -2.4052
1.3735 | -34.0554 | -2.415
1.3796 | -34.0559 | -2.424
1.3857 | -34.0562 | -2.4332
1.3918 | -34.0568 | -2.4435
1.3979 | -34.0574 | -2.4526
1.404 | -34.0578 | -2.464
1.4101 | -34.0583 | -2.4743
1.4162 | -34.0587 | -2.485
1.4223 | -34.0592 | -2.4964
1.4284 | -34.0597 | -2.5064
1.4345 | -34.0599 | -2.5181
1.4406 | -34.0602 | -2.5317
1.4467 | -34.0607 | -2.5434
1.4528 | -34.0611 | -2.5549
1.4589 | -34.0616 | -2.5668
1.465 | -34.0618 | -2.5777
1.4711 | -34.0623 | -2.5891
1.4772 | -34.0629 | -2.6026
1.4833 | -34.0633 | -2.6158
1.4894 | -34.0639 | -2.6281
1.4955 | -34.0644 | -2.6402
1.5016 | -34.065 | -2.6528
1.5078 | -34.0654 | -2.6669
1.5139 | -34.0658 | -2.6801
1.52 | -34.0662 | -2.6929
1.5261 | -34.0665 | -2.7056
1.5322 | -34.0669 | -2.7198
1.5383 | -34.0673 | -2.7325
1.5444 | -34.0678 | -2.7452
1.5505 | -34.0683 | -2.7588
1.5566 | -34.0688 | -2.7719
1.5627 | -34.0693 | -2.7859
1.5688 | -34.0699 | -2.7996
1.5749 | -34.0703 | -2.8123
1.581 | -34.0706 | -2.8258
1.5871 | -34.071 | -2.8398
1.5932 | -34.0714 | -2.8527
1.5993 | -34.0718 | -2.8672
1.6054 | -34.0721 | -2.8805
1.6115 | -34.0726 | -2.8949
1.6176 | -34.0729 | -2.9075
1.6237 | -34.0731 | -2.9213
1.6298 | -34.0733 | -2.9339
1.6359 | -34.0737 | -2.9468
1.642 | -34.0741 | -2.9588
1.6482 | -34.0743 | -2.9719
1.6543 | -34.0747 | -2.9858
1.6604 | -34.075 | -2.9979
1.6665 | -34.0751 | -3.0104
1.6726 | -34.0753 | -3.023
1.6787 | -34.0757 | -3.0356
1.6848 | -34.0756 | -3.0483
1.6909 | -34.0758 | -3.0611
1.697 | -34.076 | -3.0734
1.7031 | -34.0761 | -3.0861
1.7092 | -34.076 | -3.0983
1.7153 | -34.0762 | -3.1104
1.7214 | -34.0759 | -3.1227
1.7275 | -34.0758 | -3.1354
1.7336 | -34.0757 | -3.1486
1.7397 | -34.0759 | -3.161
1.7458 | -34.0759 | -3.1733
1.7519 | -34.0759 | -3.1856
1.758 | -34.0759 | -3.1976
1.7641 | -34.076 | -3.2093
1.7702 | -34.0759 | -3.2224
1.7763 | -34.0759 | -3.234
1.7824 | -34.0758 | -3.2465
1.7885 | -34.0758 | -3.2592
1.7947 | -34.0759 | -3.2713
1.8008 | -34.0758 | -3.2827
1.8069 | -34.0758 | -3.2943
1.813 | -34.0759 | -3.3066
1.8191 | -34.0758 | -3.3195
1.8252 | -34.0759 | -3.3325
1.8313 | -34.076 | -3.3449
1.8374 | -34.0761 | -3.3569
1.8435 | -34.0762 | -3.369
1.8496 | -34.0762 | -3.3823
1.8557 | -34.0764 | -3.3952
1.8618 | -34.0766 | -3.4077
1.8679 | -34.0767 | -3.4199
1.874 | -34.0768 | -3.4321
1.8801 | -34.0769 | -3.4447
1.8862 | -34.0772 | -3.4568
1.8923 | -34.0773 | -3.4692
1.8984 | -34.0774 | -3.4815
1.9045 | -34.0774 | -3.4932
1.9106 | -34.0774 | -3.5056
1.9167 | -34.0775 | -3.518
1.9228 | -34.0776 | -3.5301
1.9289 | -34.0777 | -3.5422
1.9351 | -34.0778 | -3.5532
1.9412 | -34.0781 | -3.5646
1.9473 | -34.0783 | -3.5769
1.9534 | -34.0785 | -3.5885
1.9595 | -34.0787 | -3.6005
1.9656 | -34.079 | -3.6131
1.9717 | -34.0792 | -3.625
1.9778 | -34.0795 | -3.6365
1.9839 | -34.0797 | -3.6482
1.99 | -34.0799 | -3.6595
1.9961 | -34.08 | -3.6718
2.0022 | -34.0802 | -3.6835
2.0083 | -34.0803 | -3.6949
2.0144 | -34.0805 | -3.7065
2.0205 | -34.0808 | -3.7185
2.0266 | -34.0812 | -3.7301
2.0327 | -34.0814 | -3.7419
2.0388 | -34.0818 | -3.7536
2.0449 | -34.082 | -3.7652
2.051 | -34.0823 | -3.7768
2.0571 | -34.0826 | -3.7883
2.0632 | -34.0828 | -3.8
2.0693 | -34.083 | -3.8124
2.0754 | -34.0832 | -3.824
2.0816 | -34.0836 | -3.8363
2.0877 | -34.0838 | -3.8481
2.0938 | -34.0841 | -3.8604
2.0999 | -34.0843 | -3.8724
2.106 | -34.0845 | -3.8842
2.1121 | -34.0847 | -3.8964
2.1182 | -34.085 | -3.9073
2.1243 | -34.0852 | -3.9182
2.1304 | -34.0854 | -3.9296
2.1365 | -34.0857 | -3.9415
2.1426 | -34.086 | -3.9537
2.1487 | -34.0863 | -3.9651
2.1548 | -34.0867 | -3.9772
2.1609 | -34.0872 | -3.9888
2.167 | -34.0876 | -4.0002
2.1731 | -34.0879 | -4.0125
2.1792 | -34.0881 | -4.0237
2.1853 | -34.0885 | -4.0359
2.1914 | -34.0888 | -4.0475
2.1975 | -34.0892 | -4.0592
2.2036 | -34.0894 | -4.0716
2.2097 | -34.0897 | -4.0833
2.2158 | -34.0901 | -4.0941
2.222 | -34.0904 | -4.1047
2.2281 | -34.0906 | -4.1163
2.2342 | -34.091 | -4.1279
2.2403 | -34.0911 | -4.1387
2.2464 | -34.0914 | -4.1496
2.2525 | -34.0917 | -4.1612
2.2586 | -34.0921 | -4.1719
2.2647 | -34.0924 | -4.1825
2.2708 | -34.0927 | -4.1939
2.2769 | -34.0931 | -4.2047
2.283 | -34.0933 | -4.216
2.2891 | -34.0935 | -4.2271
2.2952 | -34.0937 | -4.2378
2.3013 | -34.0939 | -4.2489
2.3074 | -34.0942 | -4.2594
2.3135 | -34.0944 | -4.271
2.3196 | -34.0946 | -4.2813
2.3257 | -34.0949 | -4.2922
2.3318 | -34.095 | -4.3035
2.3379 | -34.0952 | -4.3144
2.344 | -34.0953 | -4.326
2.3501 | -34.0956 | -4.3368
2.3562 | -34.0957 | -4.3474
2.3623 | -34.0959 | -4.3579
2.3685 | -34.096 | -4.3685
2.3746 | -34.0961 | -4.38
2.3807 | -34.0963 | -4.3906
2.3868 | -34.0965 | -4.4015
2.3929 | -34.0967 | -4.4123
2.399 | -34.0968 | -4.4229
2.4051 | -34.0971 | -4.4336
2.4112 | -34.0973 | -4.4444
2.4173 | -34.0975 | -4.4552
2.4234 | -34.0975 | -4.4661
2.4295 | -34.0977 | -4.4771
2.4356 | -34.0978 | -4.4881
2.4417 | -34.098 | -4.4994
2.4478 | -34.0981 | -4.5105
2.4539 | -34.0983 | -4.5208
2.46 | -34.0984 | -4.5324
2.4661 | -34.0985 | -4.5437
2.4722 | -34.0987 | -4.554
2.4783 | -34.0988 | -4.5645
2.4844 | -34.0988 | -4.5757
2.4905 | -34.0989 | -4.5869
2.4966 | -34.0991 | -4.5978
2.5027 | -34.0992 | -4.6084
2.5089 | -34.0993 | -4.6194
2.515 | -34.0993 | -4.6304
2.5211 | -34.0994 | -4.6413
2.5272 | -34.0995 | -4.6523
2.5333 | -34.0996 | -4.6634
2.5394 | -34.0998 | -4.6743
2.5455 | -34.0999 | -4.6859
2.5516 | -34.1 | -4.6963
2.5577 | -34.1002 | -4.7079
2.5638 | -34.1004 | -4.7198
2.5699 | -34.1004 | -4.7314
2.576 | -34.1004 | -4.7432
2.5821 | -34.1005 | -4.7546
2.5882 | -34.1006 | -4.7658
2.5943 | -34.1008 | -4.7768
2.6004 | -34.101 | -4.7879
2.6065 | -34.1011 | -4.7984
2.6126 | -34.1012 | -4.809
2.6187 | -34.1013 | -4.8196
2.6248 | -34.1014 | -4.8304
2.6309 | -34.1016 | -4.842
2.637 | -34.1018 | -4.8538
2.6431 | -34.102 | -4.8644
2.6492 | -34.1021 | -4.8759
2.6554 | -34.1023 | -4.8869
2.6615 | -34.1024 | -4.898
2.6676 | -34.1026 | -4.9096
2.6737 | -34.1029 | -4.9209
2.6798 | -34.1032 | -4.9321
2.6859 | -34.1033 | -4.9435
2.692 | -34.1035 | -4.9545
2.6981 | -34.1037 | -4.9658
2.7042 | -34.1039 | -4.9771
2.7103 | -34.1041 | -4.9878
2.7164 | -34.1044 | -4.9989
2.7225 | -34.1046 | -5.0104
2.7286 | -34.1048 | -5.0216
2.7347 | -34.105 | -5.0322
2.7408 | -34.1053 | -5.0424
2.7469 | -34.1055 | -5.0539
2.753 | -34.1058 | -5.0645
2.7591 | -34.1061 | -5.076
2.7652 | -34.1064 | -5.0876
2.7713 | -34.1065 | -5.0985
2.7774 | -34.1068 | -5.1096
2.7835 | -34.1069 | -5.1207
2.7896 | -34.1071 | -5.1313
2.7958 | -34.1074 | -5.1422
2.8019 | -34.1077 | -5.1527
2.808 | -34.108 | -5.1636
2.8141 | -34.1083 | -5.1749
2.8202 | -34.1086 | -5.1859
2.8263 | -34.1089 | -5.1969
2.8324 | -34.1093 | -5.2085
2.8385 | -34.1097 | -5.2194
2.8446 | -34.11 | -5.2303
2.8507 | -34.1104 | -5.2413
2.8568 | -34.1107 | -5.2519
2.8629 | -34.111 | -5.2626
2.869 | -34.1113 | -5.2739
2.8751 | -34.1117 | -5.2847
2.8812 | -34.1121 | -5.2962
2.8873 | -34.1124 | -5.3072
2.8934 | -34.1127 | -5.3184
2.8995 | -34.113 | -5.3297
2.9056 | -34.1133 | -5.341
2.9117 | -34.1137 | -5.3522
2.9178 | -34.114 | -5.3631
2.9239 | -34.1144 | -5.3748
2.93 | -34.1149 | -5.3864
2.9361 | -34.1152 | -5.3974
2.9423 | -34.1155 | -5.4082
2.9484 | -34.1159 | -5.4195
2.9545 | -34.1163 | -5.4306
2.9606 | -34.1167 | -5.4422
2.9667 | -34.1171 | -5.4535
2.9728 | -34.1174 | -5.465
2.9789 | -34.1179 | -5.4763
2.985 | -34.1183 | -5.4878
2.9911 | -34.1187 | -5.4993
2.9972 | -34.1192 | -5.5105
3.0033 | -34.1196 | -5.5223
3.0094 | -34.12 | -5.5335
3.0155 | -34.1204 | -5.5446
3.0216 | -34.1208 | -5.5561
3.0277 | -34.1211 | -5.568
3.0338 | -34.1216 | -5.5796
3.0399 | -34.122 | -5.5911
3.046 | -34.1223 | -5.6029
3.0521 | -34.1227 | -5.6141
3.0582 | -34.1229 | -5.6265
3.0643 | -34.1233 | -5.6382
3.0704 | -34.1236 | -5.6502
3.0765 | -34.1239 | -5.6615
3.0827 | -34.1242 | -5.6729
3.0888 | -34.1246 | -5.6844
3.0949 | -34.125 | -5.696
3.101 | -34.1254 | -5.7077
3.1071 | -34.1257 | -5.7197
3.1132 | -34.1261 | -5.7317
3.1193 | -34.1264 | -5.7433
3.1254 | -34.1269 | -5.7549
3.1315 | -34.1272 | -5.7671
3.1376 | -34.1275 | -5.7787
3.1437 | -34.1279 | -5.7905
3.1498 | -34.1282 | -5.8028
3.1559 | -34.1285 | -5.8145
3.162 | -34.1288 | -5.8268
3.1681 | -34.1291 | -5.839
3.1742 | -34.1294 | -5.8509
3.1803 | -34.1298 | -5.8627
3.1864 | -34.1301 | -5.8743
3.1925 | -34.1305 | -5.8861
3.1986 | -34.1307 | -5.8983
3.2047 | -34.131 | -5.9106
3.2108 | -34.1313 | -5.9219
3.2169 | -34.1316 | -5.9339
3.223 | -34.1318 | -5.9459
3.2292 | -34.1321 | -5.9579
3.2353 | -34.1323 | -5.969
3.2414 | -34.1325 | -5.9807
3.2475 | -34.1328 | -5.9923
3.2536 | -34.1331 | -6.004
3.2597 | -34.1333 | -6.0151
3.2658 | -34.1335 | -6.0263
3.2719 | -34.1338 | -6.0375
3.278 | -34.134 | -6.0493
3.2841 | -34.1343 | -6.0606
3.2902 | -34.1345 | -6.0723
3.2963 | -34.1348 | -6.0841
3.3024 | -34.1349 | -6.0953
3.3085 | -34.1352 | -6.1071
3.3146 | -34.1354 | -6.1185
3.3207 | -34.1356 | -6.1302
3.3268 | -34.1359 | -6.1419
3.3329 | -34.1361 | -6.1529
3.339 | -34.1362 | -6.1639
3.3451 | -34.1364 | -6.1753
3.3512 | -34.1366 | -6.1862
3.3573 | -34.1368 | -6.1976
3.3634 | -34.137 | -6.209
3.3696 | -34.1372 | -6.2206
3.3757 | -34.1373 | -6.2316
3.3818 | -34.1376 | -6.2426
3.3879 | -34.1378 | -6.2539
3.394 | -34.1379 | -6.2655
3.4001 | -34.1381 | -6.2765
3.4062 | -34.1382 | -6.2876
3.4123 | -34.1384 | -6.2984
3.4184 | -34.1386 | -6.3099
3.4245 | -34.1388 | -6.3209
3.4306 | -34.1389 | -6.3318
3.4367 | -34.139 | -6.3425
3.4428 | -34.1392 | -6.353
3.4489 | -34.1393 | -6.3642
3.455 | -34.1395 | -6.3752
3.4611 | -34.1397 | -6.3867
3.4672 | -34.1398 | -6.3979
3.4733 | -34.14 | -6.4092
3.4794 | -34.1402 | -6.4202
3.4855 | -34.1403 | -6.4308
3.4916 | -34.1405 | -6.4413
3.4977 | -34.1406 | -6.4528
3.5038 | -34.1408 | -6.4639
3.5099 | -34.1409 | -6.4743
3.5161 | -34.1411 | -6.4849
3.5222 | -34.1412 | -6.4954
3.5283 | -34.1413 | -6.5064
3.5344 | -34.1414 | -6.5173
3.5405 | -34.1416 | -6.5281
3.5466 | -34.1418 | -6.5392
3.5527 | -34.1419 | -6.5501
3.5588 | -34.142 | -6.5609
3.5649 | -34.142 | -6.5718
3.571 | -34.1421 | -6.5826
3.5771 | -34.1423 | -6.5939
3.5832 | -34.1424 | -6.6039
3.5893 | -34.1427 | -6.6147
3.5954 | -34.1428 | -6.6257
3.6015 | -34.143 | -6.6368
3.6076 | -34.1432 | -6.6475
3.6137 | -34.1433 | -6.6577
3.6198 | -34.1436 | -6.6681
3.6259 | -34.1437 | -6.6792
3.632 | -34.1439 | -6.69
3.6381 | -34.1441 | -6.701
3.6442 | -34.1442 | -6.7117
3.6503 | -34.1445 | -6.722
3.6565 | -34.1447 | -6.733
3.6626 | -34.1448 | -6.7437
3.6687 | -34.145 | -6.7546
3.6748 | -34.1453 | -6.7652
3.6809 | -34.1455 | -6.776
3.687 | -34.1457 | -6.7872
3.6931 | -34.146 | -6.7976
3.6992 | -34.1462 | -6.8084
3.7053 | -34.1464 | -6.819
3.7114 | -34.1465 | -6.8299
3.7175 | -34.1468 | -6.8408
3.7236 | -34.147 | -6.8509
3.7297 | -34.1472 | -6.8619
3.7358 | -34.1475 | -6.873
3.7419 | -34.1477 | -6.8841
3.748 | -34.1479 | -6.8948
3.7541 | -34.1481 | -6.9056
3.7602 | -34.1484 | -6.9162
3.7663 | -34.1486 | -6.9268
3.7724 | -34.1489 | -6.9377
3.7785 | -34.1491 | -6.9482
3.7846 | -34.1493 | -6.9588
3.7907 | -34.1497 | -6.97
3.7969 | -34.1499 | -6.9809
3.803 | -34.1502 | -6.9912
3.8091 | -34.1504 | -7.002
3.8152 | -34.1507 | -7.0126
3.8213 | -34.1509 | -7.0237
3.8274 | -34.1512 | -7.0341
3.8335 | -34.1516 | -7.0451
3.8396 | -34.1519 | -7.0559
3.8457 | -34.1522 | -7.0668
3.8518 | -34.1525 | -7.0775
3.8579 | -34.1529 | -7.0882
3.864 | -34.1532 | -7.0992
3.8701 | -34.1535 | -7.1104
3.8762 | -34.1538 | -7.121
3.8823 | -34.154 | -7.132
3.8884 | -34.1543 | -7.1429
3.8945 | -34.1547 | -7.154
3.9006 | -34.155 | -7.1652
3.9067 | -34.1553 | -7.1765
3.9128 | -34.1556 | -7.1875
3.9189 | -34.1559 | -7.1981
3.925 | -34.1562 | -7.2091
3.9311 | -34.1565 | -7.2194
3.9372 | -34.1568 | -7.2305
3.9434 | -34.1571 | -7.2414
3.9495 | -34.1575 | -7.2524
3.9556 | -34.1579 | -7.2631
3.9617 | -34.1582 | -7.2736
3.9678 | -34.1585 | -7.2842
3.9739 | -34.1588 | -7.295
3.98 | -34.1591 | -7.306
3.9861 | -34.1594 | -7.3168
3.9922 | -34.1596 | -7.3276
3.9983 | -34.1599 | -7.3381
4.0044 | -34.1603 | -7.3485
4.0105 | -34.1606 | -7.3594
4.0166 | -34.1609 | -7.3703
4.0227 | -34.1612 | -7.3811
4.0288 | -34.1615 | -7.3924
4.0349 | -34.1617 | -7.403
4.041 | -34.162 | -7.4138
4.0471 | -34.1622 | -7.4248
4.0532 | -34.1626 | -7.4352
4.0593 | -34.1629 | -7.4463
4.0654 | -34.1632 | -7.4568
4.0715 | -34.1635 | -7.4682
4.0776 | -34.1639 | -7.4789
4.0838 | -34.1642 | -7.4902
4.0899 | -34.1645 | -7.5007
4.096 | -34.1648 | -7.5115
4.1021 | -34.165 | -7.5221
4.1082 | -34.1653 | -7.5326
4.1143 | -34.1656 | -7.5431
4.1204 | -34.1659 | -7.5536
4.1265 | -34.1662 | -7.5641
4.1326 | -34.1665 | -7.5745
4.1387 | -34.1668 | -7.5851
4.1448 | -34.1671 | -7.596
4.1509 | -34.1673 | -7.6067
4.157 | -34.1676 | -7.6179
4.1631 | -34.1679 | -7.6289
4.1692 | -34.1682 | -7.6395
4.1753 | -34.1685 | -7.6502
4.1814 | -34.1687 | -7.6605
4.1875 | -34.169 | -7.6716
4.1936 | -34.1693 | -7.6824
4.1997 | -34.1695 | -7.6928
4.2058 | -34.1698 | -7.7038
4.2119 | -34.17 | -7.7143
4.218 | -34.1703 | -7.7248
4.2241 | -34.1705 | -7.7352
4.2303 | -34.1708 | -7.7457
4.2364 | -34.1711 | -7.7564
4.2425 | -34.1713 | -7.7671
4.2486 | -34.1715 | -7.7779
4.2547 | -34.1717 | -7.7883
4.2608 | -34.1719 | -7.7991
4.2669 | -34.1721 | -7.8094
4.273 | -34.1724 | -7.8202
4.2791 | -34.1726 | -7.8308
4.2852 | -34.1729 | -7.8416
4.2913 | -34.173 | -7.852
4.2974 | -34.1733 | -7.8629
4.3035 | -34.1735 | -7.8738
4.3096 | -34.1737 | -7.8846
4.3157 | -34.174 | -7.8951
4.3218 | -34.1742 | -7.9056
4.3279 | -34.1744 | -7.9162
4.334 | -34.1746 | -7.9268
4.3401 | -34.1748 | -7.9369
4.3462 | -34.1751 | -7.9476
4.3523 | -34.1754 | -7.9578
4.3584 | -34.1756 | -7.9679
4.3645 | -34.1758 | -7.9784
4.3707 | -34.176 | -7.9886
4.3768 | -34.1762 | -7.9991
4.3829 | -34.1765 | -8.01
4.389 | -34.1767 | -8.0205
4.3951 | -34.1769 | -8.0308
4.4012 | -34.1771 | -8.0417
4.4073 | -34.1773 | -8.052
4.4134 | -34.1775 | -8.0624
4.4195 | -34.1778 | -8.0729
4.4256 | -34.1781 | -8.0832
4.4317 | -34.1783 | -8.0936
4.4378 | -34.1785 | -8.1037
4.4439 | -34.1787 | -8.1139
4.45 | -34.1788 | -8.1243
4.4561 | -34.179 | -8.1348
4.4622 | -34.1792 | -8.1453
4.4683 | -34.1794 | -8.1557
4.4744 | -34.1796 | -8.1665
4.4805 | -34.1798 | -8.1765
4.4866 | -34.18 | -8.1869
4.4927 | -34.1803 | -8.1976
4.4988 | -34.1805 | -8.208
4.5049 | -34.1807 | -8.218
4.511 | -34.181 | -8.2281
4.5172 | -34.1812 | -8.2385
4.5233 | -34.1814 | -8.2493
4.5294 | -34.1816 | -8.2595
4.5355 | -34.1818 | -8.2699
4.5416 | -34.182 | -8.2803
4.5477 | -34.1823 | -8.2907
4.5538 | -34.1824 | -8.3006
4.5599 | -34.1827 | -8.3109
4.566 | -34.1829 | -8.3213
4.5721 | -34.1831 | -8.3317
4.5782 | -34.1834 | -8.3423
4.5843 | -34.1836 | -8.3528
4.5904 | -34.1839 | -8.3633
4.5965 | -34.1841 | -8.3735
4.6026 | -34.1843 | -8.3844
4.6087 | -34.1845 | -8.3946
4.6148 | -34.1847 | -8.4054
4.6209 | -34.185 | -8.4157
4.627 | -34.1852 | -8.4263
4.6331 | -34.1854 | -8.4366
4.6392 | -34.1856 | -8.4471
4.6453 | -34.1858 | -8.4574
4.6514 | -34.1861 | -8.4679
4.6576 | -34.1863 | -8.4786
4.6637 | -34.1865 | -8.4887
4.6698 | -34.1866 | -8.4989
4.6759 | -34.1869 | -8.5097
4.682 | -34.1871 | -8.5199
4.6881 | -34.1872 | -8.5303
4.6942 | -34.1875 | -8.5407
4.7003 | -34.1877 | -8.5511
4.7064 | -34.1879 | -8.5613
4.7125 | -34.1882 | -8.572
4.7186 | -34.1884 | -8.5821
4.7247 | -34.1886 | -8.5926
4.7308 | -34.1889 | -8.6029
4.7369 | -34.1891 | -8.6134
4.743 | -34.1893 | -8.6242
4.7491 | -34.1896 | -8.6349
4.7552 | -34.1898 | -8.6455
4.7613 | -34.1899 | -8.656
4.7674 | -34.1901 | -8.6663
4.7735 | -34.1904 | -8.6772
4.7796 | -34.1907 | -8.6879
4.7857 | -34.1909 | -8.6985
4.7918 | -34.1911 | -8.7087
4.7979 | -34.1914 | -8.7191
4.8041 | -34.1916 | -8.7295
4.8102 | -34.1918 | -8.7397
4.8163 | -34.1921 | -8.75
4.8224 | -34.1923 | -8.7605
4.8285 | -34.1926 | -8.7706
4.8346 | -34.1928 | -8.7808
4.8407 | -34.193 | -8.7914
4.8468 | -34.1933 | -8.8021
4.8529 | -34.1936 | -8.8129
4.859 | -34.1938 | -8.8237
4.8651 | -34.194 | -8.8338
4.8712 | -34.1943 | -8.8441
4.8773 | -34.1945 | -8.8546
4.8834 | -34.1947 | -8.8651
4.8895 | -34.195 | -8.8753
4.8956 | -34.1952 | -8.8857
4.9017 | -34.1954 | -8.896
4.9078 | -34.1956 | -8.9068
4.9139 | -34.1959 | -8.9168
4.92 | -34.1962 | -8.9269
4.9261 | -34.1965 | -8.937
4.9322 | -34.1967 | -8.9475
4.9383 | -34.197 | -8.9579
4.9445 | -34.1972 | -8.9683
4.9506 | -34.1975 | -8.9786
4.9567 | -34.1977 | -8.9888
4.9628 | -34.1979 | -8.9991
4.9689 | -34.1982 | -9.0099
4.975 | -34.1984 | -9.0207
4.9811 | -34.1987 | -9.0308
4.9872 | -34.199 | -9.0412
4.9933 | -34.1993 | -9.0519
4.9994 | -34.1996 | -9.0628
5.0055 | -34.1999 | -9.0733
5.0116 | -34.2001 | -9.0842
5.0177 | -34.2005 | -9.0949
5.0238 | -34.2007 | -9.1057
5.0299 | -34.2011 | -9.1166
5.036 | -34.2014 | -9.1271
5.0421 | -34.2016 | -9.1375
5.0482 | -34.2019 | -9.148
5.0543 | -34.2021 | -9.1582
5.0604 | -34.2025 | -9.1687
5.0665 | -34.2028 | -9.1789
5.0726 | -34.203 | -9.1897
5.0787 | -34.2033 | -9.2004
5.0848 | -34.2036 | -9.211
5.091 | -34.2039 | -9.2213
5.0971 | -34.2042 | -9.2315
5.1032 | -34.2045 | -9.2423
5.1093 | -34.2048 | -9.2528
5.1154 | -34.205 | -9.263
5.1215 | -34.2053 | -9.2733
5.1276 | -34.2056 | -9.2836
5.1337 | -34.2059 | -9.2938
5.1398 | -34.2062 | -9.304
5.1459 | -34.2065 | -9.3148
5.152 | -34.2067 | -9.3254
5.1581 | -34.207 | -9.3361
5.1642 | -34.2073 | -9.3467
5.1703 | -34.2076 | -9.3572
5.1764 | -34.208 | -9.3676
5.1825 | -34.2082 | -9.3778
5.1886 | -34.2085 | -9.388
5.1947 | -34.2088 | -9.3985
5.2008 | -34.2091 | -9.4091
5.2069 | -34.2094 | -9.4199
5.213 | -34.2097 | -9.4302
5.2191 | -34.21 | -9.4407
5.2252 | -34.2103 | -9.4513
5.2314 | -34.2106 | -9.462
5.2375 | -34.2109 | -9.4722
5.2436 | -34.2111 | -9.4828
5.2497 | -34.2115 | -9.493
5.2558 | -34.2118 | -9.5039
5.2619 | -34.2121 | -9.5147
5.268 | -34.2124 | -9.5254
5.2741 | -34.2127 | -9.536
5.2802 | -34.213 | -9.547
5.2863 | -34.2134 | -9.5574
5.2924 | -34.2137 | -9.5682
5.2985 | -34.214 | -9.5788
5.3046 | -34.2143 | -9.5894
5.3107 | -34.2146 | -9.6002
5.3168 | -34.2149 | -9.6108
5.3229 | -34.2152 | -9.6214
5.329 | -34.2155 | -9.632
5.3351 | -34.2159 | -9.6422
5.3412 | -34.2162 | -9.6528
5.3473 | -34.2165 | -9.6634
5.3534 | -34.2168 | -9.6741
5.3595 | -34.2171 | -9.6848
5.3656 | -34.2175 | -9.6953
5.3717 | -34.2177 | -9.7063
5.3779 | -34.218 | -9.7172
5.384 | -34.2184 | -9.7278
5.3901 | -34.2187 | -9.7387
5.3962 | -34.219 | -9.7491
5.4023 | -34.2193 | -9.7599
5.4084 | -34.2197 | -9.7705
5.4145 | -34.22 | -9.7808
5.4206 | -34.2203 | -9.7914
5.4267 | -34.2206 | -9.8024
5.4328 | -34.2209 | -9.8131
5.4389 | -34.2213 | -9.8234
5.445 | -34.2215 | -9.8344
5.4511 | -34.2219 | -9.8448
5.4572 | -34.2222 | -9.8555
5.4633 | -34.2225 | -9.8662
5.4694 | -34.2228 | -9.8768
5.4755 | -34.2231 | -9.8872
5.4816 | -34.2235 | -9.8978
5.4877 | -34.2238 | -9.9083
5.4938 | -34.2241 | -9.9187
5.4999 | -34.2244 | -9.9292
5.506 | -34.2248 | -9.9398
5.5121 | -34.2251 | -9.9503
5.5183 | -34.2254 | -9.9608
5.5244 | -34.2257 | -9.9716
5.5305 | -34.2261 | -9.9821
5.5366 | -34.2264 | -9.9925
5.5427 | -34.2266 | -10.0032
5.5488 | -34.227 | -10.0139
5.5549 | -34.2273 | -10.0246
5.561 | -34.2276 | -10.0355
5.5671 | -34.2279 | -10.0459
5.5732 | -34.2282 | -10.0568
5.5793 | -34.2286 | -10.0678
5.5854 | -34.229 | -10.0784
5.5915 | -34.2292 | -10.0893
5.5976 | -34.2296 | -10.1
5.6037 | -34.2299 | -10.1107
5.6098 | -34.2302 | -10.1213
5.6159 | -34.2305 | -10.1318
5.622 | -34.2308 | -10.1425
5.6281 | -34.2311 | -10.1533
5.6342 | -34.2315 | -10.1643
5.6403 | -34.2318 | -10.1749
5.6464 | -34.2321 | -10.1858
5.6525 | -34.2324 | -10.1965
5.6586 | -34.2328 | -10.2068
5.6648 | -34.2331 | -10.2175
5.6709 | -34.2335 | -10.2284
5.677 | -34.2338 | -10.2391
5.6831 | -34.2341 | -10.2498
5.6892 | -34.2344 | -10.2607
5.6953 | -34.2347 | -10.2713
5.7014 | -34.2351 | -10.2821
5.7075 | -34.2353 | -10.2927
5.7136 | -34.2357 | -10.3034
5.7197 | -34.236 | -10.314
5.7258 | -34.2363 | -10.3245
5.7319 | -34.2366 | -10.3352
5.738 | -34.237 | -10.3458
5.7441 | -34.2373 | -10.3565
5.7502 | -34.2377 | -10.3668
5.7563 | -34.238 | -10.3775
5.7624 | -34.2383 | -10.3883
5.7685 | -34.2387 | -10.3991
5.7746 | -34.239 | -10.41
5.7807 | -34.2392 | -10.4207
5.7868 | -34.2395 | -10.4316
5.7929 | -34.2398 | -10.4422
5.799 | -34.2402 | -10.4527
5.8052 | -34.2405 | -10.4632
5.8113 | -34.2408 | -10.4738
5.8174 | -34.2411 | -10.4843
5.8235 | -34.2414 | -10.4949
5.8296 | -34.2417 | -10.5057
5.8357 | -34.242 | -10.5163
5.8418 | -34.2424 | -10.5271
5.8479 | -34.2426 | -10.5379
5.854 | -34.243 | -10.5487
5.8601 | -34.2433 | -10.5592
5.8662 | -34.2436 | -10.57
5.8723 | -34.2439 | -10.5804
5.8784 | -34.2442 | -10.591
5.8845 | -34.2445 | -10.6015
5.8906 | -34.2448 | -10.6122
5.8967 | -34.2451 | -10.6227
5.9028 | -34.2455 | -10.6331
5.9089 | -34.2459 | -10.6436
5.915 | -34.2462 | -10.654
5.9211 | -34.2465 | -10.6646
5.9272 | -34.2468 | -10.6749
5.9333 | -34.2471 | -10.6854
5.9394 | -34.2474 | -10.6961
5.9455 | -34.2478 | -10.7064
5.9517 | -34.2481 | -10.7168
5.9578 | -34.2485 | -10.7276
5.9639 | -34.2488 | -10.7378
5.97 | -34.2492 | -10.7485
5.9761 | -34.2495 | -10.759
5.9822 | -34.2498 | -10.7696
5.9883 | -34.2502 | -10.7802
5.9944 | -34.2505 | -10.7909
6.0005 | -34.2508 | -10.8016
6.0066 | -34.2511 | -10.812
6.0127 | -34.2515 | -10.8223
6.0188 | -34.2518 | -10.8332
6.0249 | -34.2521 | -10.8438
6.031 | -34.2524 | -10.8542
6.0371 | -34.2528 | -10.8648
6.0432 | -34.2531 | -10.8752
6.0493 | -34.2534 | -10.8856
6.0554 | -34.2538 | -10.896
6.0615 | -34.2541 | -10.9065
6.0676 | -34.2545 | -10.9173
6.0737 | -34.2548 | -10.928
6.0798 | -34.2551 | -10.9385
6.0859 | -34.2555 | -10.9489
6.0921 | -34.2558 | -10.9592
6.0982 | -34.2561 | -10.9697
6.1043 | -34.2565 | -10.9801
6.1104 | -34.2568 | -10.9909
6.1165 | -34.2572 | -11.0009
6.1226 | -34.2575 | -11.0113
6.1287 | -34.2578 | -11.0218
6.1348 | -34.2582 | -11.0323
6.1409 | -34.2585 | -11.0427
6.147 | -34.2588 | -11.0531
6.1531 | -34.2591 | -11.0636
6.1592 | -34.2595 | -11.0739
6.1653 | -34.2598 | -11.0843
6.1714 | -34.2601 | -11.0948
6.1775 | -34.2605 | -11.1054
6.1836 | -34.2608 | -11.1159
6.1897 | -34.2611 | -11.1263
6.1958 | -34.2614 | -11.1371
6.2019 | -34.2618 | -11.1477
6.208 | -34.2622 | -11.1582
6.2141 | -34.2625 | -11.1685
6.2202 | -34.2629 | -11.1791
6.2263 | -34.2632 | -11.1895
6.2325 | -34.2635 | -11.1996
6.2386 | -34.2638 | -11.2097
6.2447 | -34.2642 | -11.2205
6.2508 | -34.2645 | -11.2309
6.2569 | -34.2649 | -11.2413
6.263 | -34.2652 | -11.2517
6.2691 | -34.2655 | -11.2621
6.2752 | -34.2659 | -11.2726
6.2813 | -34.2662 | -11.283
6.2874 | -34.2665 | -11.2933
6.2935 | -34.2668 | -11.3036
6.2996 | -34.2671 | -11.314
6.3057 | -34.2675 | -11.3245
6.3118 | -34.2677 | -11.335
6.3179 | -34.2681 | -11.3454
6.324 | -34.2684 | -11.3558
6.3301 | -34.2688 | -11.3661
6.3362 | -34.2691 | -11.3764
6.3423 | -34.2695 | -11.387
6.3484 | -34.2698 | -11.3975
6.3545 | -34.2701 | -11.4075
6.3606 | -34.2704 | -11.4176
6.3667 | -34.2707 | -11.4279
6.3728 | -34.2711 | -11.4383
6.379 | -34.2714 | -11.4487
6.3851 | -34.2718 | -11.4588
6.3912 | -34.272 | -11.4691
6.3973 | -34.2723 | -11.4795
6.4034 | -34.2727 | -11.4897
6.4095 | -34.273 | -11.5
6.4156 | -34.2733 | -11.5104
6.4217 | -34.2736 | -11.5208
6.4278 | -34.274 | -11.5312
6.4339 | -34.2743 | -11.5415
6.44 | -34.2747 | -11.5517
6.4461 | -34.275 | -11.5623
6.4522 | -34.2753 | -11.5724
6.4583 | -34.2756 | -11.5829
6.4644 | -34.2759 | -11.5933
6.4705 | -34.2762 | -11.6033
6.4766 | -34.2766 | -11.6136
6.4827 | -34.2769 | -11.624
6.4888 | -34.2772 | -11.6341
6.4949 | -34.2775 | -11.6441
6.501 | -34.2779 | -11.6545
6.5071 | -34.2782 | -11.6646
6.5132 | -34.2785 | -11.6749
6.5194 | -34.2788 | -11.6853
6.5255 | -34.2792 | -11.6953
6.5316 | -34.2795 | -11.7058
6.5377 | -34.2798 | -11.7161
6.5438 | -34.2801 | -11.7261
6.5499 | -34.2804 | -11.7363
6.556 | -34.2808 | -11.7465
6.5621 | -34.2811 | -11.7571
6.5682 | -34.2814 | -11.7674
6.5743 | -34.2818 | -11.7776
6.5804 | -34.2821 | -11.788
6.5865 | -34.2824 | -11.7983
6.5926 | -34.2828 | -11.8084
6.5987 | -34.2831 | -11.8184
6.6048 | -34.2834 | -11.8287
6.6109 | -34.2837 | -11.8389
6.617 | -34.284 | -11.8493
6.6231 | -34.2843 | -11.8596
6.6292 | -34.2847 | -11.8698
6.6353 | -34.285 | -11.8804
6.6414 | -34.2854 | -11.8905
6.6475 | -34.2857 | -11.9007
6.6536 | -34.286 | -11.9109
6.6597 | -34.2863 | -11.9208
6.6659 | -34.2866 | -11.9311
6.672 | -34.287 | -11.9412
6.6781 | -34.2873 | -11.9514
6.6842 | -34.2876 | -11.9615
6.6903 | -34.288 | -11.9717
6.6964 | -34.2884 | -11.982
6.7025 | -34.2887 | -11.992
6.7086 | -34.289 | -12.0022
6.7147 | -34.2893 | -12.0124
6.7208 | -34.2896 | -12.0228
6.7269 | -34.2899 | -12.0329
6.733 | -34.2902 | -12.0429
6.7391 | -34.2906 | -12.0531
6.7452 | -34.2909 | -12.063
6.7513 | -34.2912 | -12.0729
6.7574 | -34.2916 | -12.0832
6.7635 | -34.2919 | -12.0932
6.7696 | -34.2923 | -12.1032
6.7757 | -34.2926 | -12.1134
6.7818 | -34.293 | -12.1236
6.7879 | -34.2933 | -12.1335
6.794 | -34.2936 | -12.1435
6.8001 | -34.2939 | -12.1536
6.8063 | -34.2943 | -12.1637
6.8124 | -34.2946 | -12.174
6.8185 | -34.2949 | -12.1844
6.8246 | -34.2952 | -12.1943
6.8307 | -34.2955 | -12.2043
6.8368 | -34.2959 | -12.2147
6.8429 | -34.2962 | -12.2249
6.849 | -34.2965 | -12.2351
6.8551 | -34.2969 | -12.2451
6.8612 | -34.2972 | -12.2552
6.8673 | -34.2975 | -12.2652
6.8734 | -34.2979 | -12.2755
6.8795 | -34.2982 | -12.2857
6.8856 | -34.2986 | -12.2958
6.8917 | -34.2989 | -12.3058
6.8978 | -34.2992 | -12.3157
6.9039 | -34.2995 | -12.3255
6.91 | -34.2999 | -12.3356
6.9161 | -34.3002 | -12.3455
6.9222 | -34.3005 | -12.3556
6.9283 | -34.3009 | -12.3654
6.9344 | -34.3013 | -12.3755
6.9405 | -34.3016 | -12.3856
6.9466 | -34.302 | -12.3959
6.9528 | -34.3023 | -12.4062
6.9589 | -34.3027 | -12.4164
6.965 | -34.303 | -12.4266
6.9711 | -34.3034 | -12.4368
6.9772 | -34.3038 | -12.4469
6.9833 | -34.3041 | -12.4569
6.9894 | -34.3044 | -12.4673
6.9955 | -34.3048 | -12.4771
7.0016 | -34.3051 | -12.4871
7.0077 | -34.3055 | -12.4972
7.0138 | -34.3058 | -12.5072
7.0199 | -34.3062 | -12.5172
7.026 | -34.3065 | -12.5272
7.0321 | -34.3069 | -12.537
7.0382 | -34.3072 | -12.5468
7.0443 | -34.3076 | -12.557
7.0504 | -34.3079 | -12.5672
7.0565 | -34.3083 | -12.5772
7.0626 | -34.3086 | -12.587
7.0687 | -34.309 | -12.5966
7.0748 | -34.3093 | -12.6064
7.0809 | -34.3097 | -12.6164
7.087 | -34.31 | -12.6265
7.0932 | -34.3103 | -12.6366
7.0993 | -34.3107 | -12.6465
7.1054 | -34.3111 | -12.6564
7.1115 | -34.3114 | -12.6662
7.1176 | -34.3118 | -12.6761
7.1237 | -34.3121 | -12.6861
7.1298 | -34.3124 | -12.696
7.1359 | -34.3128 | -12.706
7.142 | -34.3131 | -12.7157
7.1481 | -34.3135 | -12.7256
7.1542 | -34.3138 | -12.7356
7.1603 | -34.3142 | -12.7456
7.1664 | -34.3146 | -12.7555
7.1725 | -34.3149 | -12.7655
7.1786 | -34.3153 | -12.7755
7.1847 | -34.3156 | -12.7855
7.1908 | -34.316 | -12.7957
7.1969 | -34.3163 | -12.8058
7.203 | -34.3167 | -12.8156
7.2091 | -34.317 | -12.8255
7.2152 | -34.3174 | -12.8355
7.2213 | -34.3177 | -12.8453
7.2274 | -34.3181 | -12.8552
7.2335 | -34.3184 | -12.8651
7.2397 | -34.3188 | -12.875
7.2458 | -34.3192 | -12.8847
7.2519 | -34.3196 | -12.8947
7.258 | -34.3199 | -12.9048
7.2641 | -34.3203 | -12.9147
7.2702 | -34.3206 | -12.9249
7.2763 | -34.321 | -12.9349
7.2824 | -34.3214 | -12.945
7.2885 | -34.3217 | -12.955
7.2946 | -34.3221 | -12.965
7.3007 | -34.3224 | -12.975
7.3068 | -34.3227 | -12.9853
7.3129 | -34.3231 | -12.9954
7.319 | -34.3235 | -13.0054
7.3251 | -34.3238 | -13.0151
7.3312 | -34.3242 | -13.0251
7.3373 | -34.3246 | -13.0348
7.3434 | -34.3249 | -13.0446
7.3495 | -34.3253 | -13.0545
7.3556 | -34.3257 | -13.0645
7.3617 | -34.3261 | -13.0747
7.3678 | -34.3264 | -13.0846
7.3739 | -34.3267 | -13.0947
7.3801 | -34.3271 | -13.105
7.3862 | -34.3275 | -13.1149
7.3923 | -34.3279 | -13.125
7.3984 | -34.3282 | -13.1347
7.4045 | -34.3286 | -13.1446
7.4106 | -34.329 | -13.1545
7.4167 | -34.3293 | -13.1645
7.4228 | -34.3297 | -13.1746
7.4289 | -34.3301 | -13.1845
7.435 | -34.3305 | -13.1943
7.4411 | -34.3308 | -13.2044
7.4472 | -34.3312 | -13.2144
7.4533 | -34.3316 | -13.2241
7.4594 | -34.332 | -13.2344
7.4655 | -34.3323 | -13.2441
7.4716 | -34.3327 | -13.2543
7.4777 | -34.333 | -13.2642
7.4838 | -34.3334 | -13.274
7.4899 | -34.3338 | -13.2839
7.496 | -34.3341 | -13.2937
7.5021 | -34.3345 | -13.3036
7.5082 | -34.3349 | -13.3136
7.5143 | -34.3352 | -13.3236
7.5204 | -34.3356 | -13.3337
7.5266 | -34.336 | -13.3435
7.5327 | -34.3363 | -13.3533
7.5388 | -34.3367 | -13.3636
7.5449 | -34.3371 | -13.3738
7.551 | -34.3375 | -13.3836
7.5571 | -34.3378 | -13.3937
7.5632 | -34.3382 | -13.4037
7.5693 | -34.3386 | -13.4136
7.5754 | -34.339 | -13.4234
7.5815 | -34.3393 | -13.4334
7.5876 | -34.3397 | -13.4436
7.5937 | -34.3401 | -13.4537
7.5998 | -34.3405 | -13.4636
7.6059 | -34.3409 | -13.4738
7.612 | -34.3413 | -13.4837
7.6181 | -34.3416 | -13.4936
7.6242 | -34.342 | -13.5039
7.6303 | -34.3424 | -13.5141
7.6364 | -34.3427 | -13.5242
7.6425 | -34.3431 | -13.5342
7.6486 | -34.3435 | -13.5442
7.6547 | -34.3439 | -13.5541
7.6608 | -34.3442 | -13.5641
7.667 | -34.3446 | -13.5742
7.6731 | -34.345 | -13.5842
7.6792 | -34.3454 | -13.5945
7.6853 | -34.3457 | -13.6046
7.6914 | -34.3461 | -13.6147
7.6975 | -34.3465 | -13.6248
7.7036 | -34.3469 | -13.6349
7.7097 | -34.3473 | -13.6451
7.7158 | -34.3477 | -13.6552
7.7219 | -34.348 | -13.6651
7.728 | -34.3484 | -13.6752
7.7341 | -34.3488 | -13.6854
7.7402 | -34.3492 | -13.6955
7.7463 | -34.3496 | -13.7056
7.7524 | -34.3499 | -13.7159
7.7585 | -34.3503 | -13.726
7.7646 | -34.3507 | -13.736
7.7707 | -34.351 | -13.7463
7.7768 | -34.3514 | -13.7565
7.7829 | -34.3518 | -13.7668
7.789 | -34.3521 | -13.7771
7.7951 | -34.3525 | -13.7872
7.8012 | -34.3529 | -13.7972
7.8073 | -34.3533 | -13.8076
7.8135 | -34.3537 | -13.8176
7.8196 | -34.3541 | -13.828
7.8257 | -34.3545 | -13.8383
7.8318 | -34.3548 | -13.8484
7.8379 | -34.3552 | -13.8584
7.844 | -34.3556 | -13.8685
7.8501 | -34.356 | -13.8787
7.8562 | -34.3563 | -13.8891
7.8623 | -34.3568 | -13.8991
7.8684 | -34.3571 | -13.9093
7.8745 | -34.3575 | -13.9194
7.8806 | -34.3579 | -13.9297
7.8867 | -34.3583 | -13.9396
7.8928 | -34.3587 | -13.9495
7.8989 | -34.359 | -13.9596
7.905 | -34.3594 | -13.9698
7.9111 | -34.3598 | -13.98
7.9172 | -34.3602 | -13.9901
7.9233 | -34.3605 | -14.0003
7.9294 | -34.3609 | -14.0106
7.9355 | -34.3613 | -14.0205
7.9416 | -34.3616 | -14.0306
7.9477 | -34.362 | -14.0406
7.9539 | -34.3624 | -14.0507
7.96 | -34.3628 | -14.061
7.9661 | -34.3632 | -14.0711
7.9722 | -34.3636 | -14.0811
7.9783 | -34.364 | -14.0915
7.9844 | -34.3644 | -14.1016
7.9905 | -34.3648 | -14.112
7.9966 | -34.3651 | -14.1223
8.0027 | -34.3655 | -14.1324
8.0088 | -34.3659 | -14.1425
8.0149 | -34.3662 | -14.1527
8.021 | -34.3666 | -14.1629
8.0271 | -34.367 | -14.1733
8.0332 | -34.3674 | -14.1834
8.0393 | -34.3678 | -14.1937
8.0454 | -34.3682 | -14.2041
8.0515 | -34.3686 | -14.2143
8.0576 | -34.369 | -14.2247
8.0637 | -34.3693 | -14.2348
8.0698 | -34.3697 | -14.2448
8.0759 | -34.37 | -14.2548
8.082 | -34.3704 | -14.2649
8.0881 | -34.3708 | -14.2751
8.0942 | -34.3712 | -14.2854
8.1004 | -34.3717 | -14.2957
8.1065 | -34.372 | -14.3056
8.1126 | -34.3724 | -14.3158
8.1187 | -34.3728 | -14.3261
8.1248 | -34.3731 | -14.336
8.1309 | -34.3735 | -14.346
8.137 | -34.3739 | -14.356
8.1431 | -34.3743 | -14.3664
8.1492 | -34.3747 | -14.3765
8.1553 | -34.375 | -14.3867
8.1614 | -34.3754 | -14.3971
8.1675 | -34.3758 | -14.4072
8.1736 | -34.3761 | -14.4176
8.1797 | -34.3765 | -14.4278
8.1858 | -34.3769 | -14.4381
8.1919 | -34.3773 | -14.4482
8.198 | -34.3777 | -14.4582
8.2041 | -34.3781 | -14.4684
8.2102 | -34.3785 | -14.4786
8.2163 | -34.3789 | -14.4885
8.2224 | -34.3792 | -14.4987
8.2285 | -34.3796 | -14.5085
8.2346 | -34.38 | -14.5186
8.2408 | -34.3804 | -14.5286
8.2469 | -34.3808 | -14.5389
8.253 | -34.3812 | -14.5489
8.2591 | -34.3816 | -14.5591
8.2652 | -34.382 | -14.5693
8.2713 | -34.3824 | -14.5795
8.2774 | -34.3828 | -14.5898
8.2835 | -34.3832 | -14.6002
8.2896 | -34.3836 | -14.6102
8.2957 | -34.384 | -14.6204
8.3018 | -34.3844 | -14.6306
8.3079 | -34.3848 | -14.6405
8.314 | -34.3852 | -14.6507
8.3201 | -34.3856 | -14.6607
8.3262 | -34.3859 | -14.6709
8.3323 | -34.3863 | -14.6811
8.3384 | -34.3867 | -14.6914
8.3445 | -34.3871 | -14.7014
8.3506 | -34.3875 | -14.7116
8.3567 | -34.3879 | -14.7219
8.3628 | -34.3883 | -14.732
8.3689 | -34.3887 | -14.742
8.375 | -34.3891 | -14.7522
8.3812 | -34.3894 | -14.7623
8.3873 | -34.3898 | -14.7724
8.3934 | -34.3902 | -14.7827
8.3995 | -34.3906 | -14.7927
8.4056 | -34.391 | -14.8027
8.4117 | -34.3914 | -14.813
8.4178 | -34.3917 | -14.8232
8.4239 | -34.3921 | -14.8331
8.43 | -34.3925 | -14.8433
8.4361 | -34.3928 | -14.8535
8.4422 | -34.3933 | -14.8635
8.4483 | -34.3936 | -14.8735
8.4544 | -34.3941 | -14.8837
8.4605 | -34.3944 | -14.8937
8.4666 | -34.3948 | -14.9041
8.4727 | -34.3952 | -14.9142
8.4788 | -34.3956 | -14.9243
8.4849 | -34.396 | -14.9344
8.491 | -34.3964 | -14.9445
8.4971 | -34.3968 | -14.9544
8.5032 | -34.3971 | -14.9643
8.5093 | -34.3975 | -14.9743
8.5154 | -34.3979 | -14.9845
8.5215 | -34.3983 | -14.9944
8.5277 | -34.3987 | -15.0045
8.5338 | -34.3991 | -15.0146
8.5399 | -34.3995 | -15.0246
8.546 | -34.3999 | -15.0345
8.5521 | -34.4003 | -15.0446
8.5582 | -34.4006 | -15.0547
8.5643 | -34.401 | -15.0645
8.5704 | -34.4014 | -15.0742
8.5765 | -34.4018 | -15.0844
8.5826 | -34.4022 | -15.0941
8.5887 | -34.4026 | -15.1041
8.5948 | -34.403 | -15.1142
8.6009 | -34.4034 | -15.124
8.607 | -34.4038 | -15.1342
8.6131 | -34.4042 | -15.1443
8.6192 | -34.4046 | -15.1542
8.6253 | -34.405 | -15.1641
8.6314 | -34.4054 | -15.1741
8.6375 | -34.4058 | -15.1842
8.6436 | -34.4062 | -15.1941
8.6497 | -34.4066 | -15.204
8.6558 | -34.4069 | -15.214
8.6619 | -34.4073 | -15.2241
8.6681 | -34.4077 | -15.2342
8.6742 | -34.4081 | -15.2442
8.6803 | -34.4085 | -15.2542
8.6864 | -34.4089 | -15.2639
8.6925 | -34.4093 | -15.2737
8.6986 | -34.4097 | -15.2838
8.7047 | -34.41 | -15.294
8.7108 | -34.4104 | -15.304
8.7169 | -34.4108 | -15.314
8.723 | -34.4112 | -15.3237
8.7291 | -34.4116 | -15.3335
8.7352 | -34.412 | -15.3437
8.7413 | -34.4124 | -15.3538
8.7474 | -34.4128 | -15.3635
8.7535 | -34.4132 | -15.3735
8.7596 | -34.4136 | -15.3834
8.7657 | -34.414 | -15.3935
8.7718 | -34.4144 | -15.4035
8.7779 | -34.4148 | -15.4136
8.784 | -34.4152 | -15.4236
8.7901 | -34.4156 | -15.4336
8.7962 | -34.416 | -15.4433
8.8023 | -34.4164 | -15.4532
8.8084 | -34.4168 | -15.4631
8.8146 | -34.4172 | -15.473
8.8207 | -34.4175 | -15.4828
8.8268 | -34.418 | -15.4924
8.8329 | -34.4183 | -15.5023
8.839 | -34.4188 | -15.5123
8.8451 | -34.4191 | -15.5221
8.8512 | -34.4195 | -15.5319
8.8573 | -34.4199 | -15.5418
8.8634 | -34.4203 | -15.5515
8.8695 | -34.4207 | -15.5616
8.8756 | -34.4211 | -15.5715
8.8817 | -34.4215 | -15.5813
8.8878 | -34.4219 | -15.5915
8.8939 | -34.4223 | -15.6013
8.9 | -34.4227 | -15.6111
8.9061 | -34.4231 | -15.621
8.9122 | -34.4235 | -15.6309
8.9183 | -34.4239 | -15.6409
8.9244 | -34.4243 | -15.6509
8.9305 | -34.4247 | -15.6606
8.9366 | -34.4251 | -15.6705
8.9427 | -34.4255 | -15.6804
8.9488 | -34.4259 | -15.6902
8.955 | -34.4263 | -15.7001
8.9611 | -34.4267 | -15.7102
8.9672 | -34.4272 | -15.72
8.9733 | -34.4276 | -15.7297
8.9794 | -34.428 | -15.7395
8.9855 | -34.4284 | -15.7495
8.9916 | -34.4288 | -15.7595
8.9977 | -34.4292 | -15.7694
9.0038 | -34.4297 | -15.7791
9.0099 | -34.4301 | -15.789
9.016 | -34.4305 | -15.799
9.0221 | -34.4309 | -15.8087
9.0282 | -34.4313 | -15.8184
9.0343 | -34.4317 | -15.8283
9.0404 | -34.4321 | -15.8382
9.0465 | -34.4326 | -15.8482
9.0526 | -34.433 | -15.858
9.0587 | -34.4334 | -15.8678
9.0648 | -34.4339 | -15.8776
9.0709 | -34.4342 | -15.8874
9.077 | -34.4347 | -15.8975
9.0831 | -34.4351 | -15.9075
9.0892 | -34.4355 | -15.9174
9.0953 | -34.436 | -15.9273
9.1015 | -34.4364 | -15.9371
9.1076 | -34.4368 | -15.9468
9.1137 | -34.4372 | -15.9566
9.1198 | -34.4377 | -15.9664
9.1259 | -34.4381 | -15.9762
9.132 | -34.4385 | -15.9863
9.1381 | -34.4389 | -15.996
9.1442 | -34.4394 | -16.0054
9.1503 | -34.4398 | -16.0154
9.1564 | -34.4402 | -16.0252
9.1625 | -34.4406 | -16.035
9.1686 | -34.4411 | -16.0448
9.1747 | -34.4415 | -16.0546
9.1808 | -34.442 | -16.0643
9.1869 | -34.4424 | -16.0741
9.193 | -34.4428 | -16.0839
9.1991 | -34.4432 | -16.0938
9.2052 | -34.4436 | -16.1036
9.2113 | -34.4441 | -16.1134
9.2174 | -34.4445 | -16.1232
9.2235 | -34.4449 | -16.133
9.2296 | -34.4454 | -16.1428
9.2357 | -34.4458 | -16.1528
9.2419 | -34.4462 | -16.1626
9.248 | -34.4467 | -16.1723
9.2541 | -34.4472 | -16.1822
9.2602 | -34.4476 | -16.1921
9.2663 | -34.448 | -16.2016
9.2724 | -34.4484 | -16.2113
9.2785 | -34.4488 | -16.2211
9.2846 | -34.4493 | -16.2307
9.2907 | -34.4497 | -16.2405
9.2968 | -34.4501 | -16.2504
9.3029 | -34.4506 | -16.2603
9.309 | -34.451 | -16.2701
9.3151 | -34.4514 | -16.2798
9.3212 | -34.4519 | -16.2895
9.3273 | -34.4523 | -16.2992
9.3334 | -34.4527 | -16.3088
9.3395 | -34.4532 | -16.3185
9.3456 | -34.4536 | -16.328
9.3517 | -34.454 | -16.3379
9.3578 | -34.4544 | -16.3476
9.3639 | -34.4549 | -16.3573
9.37 | -34.4553 | -16.3671
9.3761 | -34.4558 | -16.377
9.3822 | -34.4562 | -16.3868
9.3884 | -34.4567 | -16.3962
9.3945 | -34.4571 | -16.4061
9.4006 | -34.4576 | -16.4159
9.4067 | -34.458 | -16.4257
9.4128 | -34.4584 | -16.4356
9.4189 | -34.4589 | -16.4455
9.425 | -34.4593 | -16.4552
9.4311 | -34.4598 | -16.4649
9.4372 | -34.4602 | -16.4747
9.4433 | -34.4607 | -16.4844
9.4494 | -34.4611 | -16.494
9.4555 | -34.4616 | -16.5038
9.4616 | -34.462 | -16.5136
9.4677 | -34.4625 | -16.5233
9.4738 | -34.4629 | -16.533
9.4799 | -34.4634 | -16.5428
9.486 | -34.4638 | -16.5526
9.4921 | -34.4642 | -16.5623
9.4982 | -34.4647 | -16.5718
9.5043 | -34.4652 | -16.5815
9.5104 | -34.4656 | -16.5911
9.5165 | -34.466 | -16.6008
9.5226 | -34.4665 | -16.6107
9.5288 | -34.4669 | -16.6202
9.5349 | -34.4673 | -16.63
9.541 | -34.4678 | -16.6396
9.5471 | -34.4682 | -16.6493
9.5532 | -34.4687 | -16.6591
9.5593 | -34.4691 | -16.6689
9.5654 | -34.4696 | -16.6785
9.5715 | -34.47 | -16.6882
9.5776 | -34.4704 | -16.6981
9.5837 | -34.4708 | -16.7077
9.5898 | -34.4713 | -16.7175
9.5959 | -34.4717 | -16.7273
9.602 | -34.4721 | -16.7368
9.6081 | -34.4726 | -16.7465
9.6142 | -34.4731 | -16.7561
9.6203 | -34.4735 | -16.7658
9.6264 | -34.4739 | -16.7754
9.6325 | -34.4743 | -16.7852
9.6386 | -34.4748 | -16.7948
9.6447 | -34.4752 | -16.8045
9.6508 | -34.4757 | -16.8141
9.6569 | -34.4761 | -16.8238
9.663 | -34.4766 | -16.8335
9.6691 | -34.477 | -16.8433
9.6753 | -34.4775 | -16.8529
9.6814 | -34.4779 | -16.8625
9.6875 | -34.4783 | -16.872
9.6936 | -34.4788 | -16.8815
9.6997 | -34.4793 | -16.8914
9.7058 | -34.4797 | -16.9009
9.7119 | -34.4802 | -16.9107
9.718 | -34.4806 | -16.9206
9.7241 | -34.4811 | -16.9301
9.7302 | -34.4815 | -16.9399
9.7363 | -34.4819 | -16.9494
9.7424 | -34.4824 | -16.9589
9.7485 | -34.4828 | -16.9687
9.7546 | -34.4833 | -16.9782
9.7607 | -34.4837 | -16.9881
9.7668 | -34.4842 | -16.9978
9.7729 | -34.4846 | -17.0075
9.779 | -34.4851 | -17.0169
9.7851 | -34.4856 | -17.0267
9.7912 | -34.486 | -17.0363
9.7973 | -34.4865 | -17.0461
9.8034 | -34.487 | -17.0556
9.8095 | -34.4874 | -17.0655
9.8157 | -34.4879 | -17.0751
9.8218 | -34.4883 | -17.085
9.8279 | -34.4888 | -17.0945
9.834 | -34.4892 | -17.1042
9.8401 | -34.4896 | -17.1139
9.8462 | -34.4901 | -17.1233
9.8523 | -34.4905 | -17.133
9.8584 | -34.491 | -17.1426
9.8645 | -34.4914 | -17.1523
9.8706 | -34.4918 | -17.162
9.8767 | -34.4923 | -17.1716
9.8828 | -34.4927 | -17.1811
9.8889 | -34.4931 | -17.1906
9.895 | -34.4936 | -17.2003
9.9011 | -34.494 | -17.2099
9.9072 | -34.4945 | -17.2195
9.9133 | -34.4949 | -17.2292
9.9194 | -34.4954 | -17.2387
9.9255 | -34.4958 | -17.2482
9.9316 | -34.4963 | -17.2578
9.9377 | -34.4967 | -17.2674
9.9438 | -34.4971 | -17.2771
9.9499 | -34.4976 | -17.287
9.956 | -34.498 | -17.2966
9.9622 | -34.4985 | -17.3062
9.9683 | -34.4989 | -17.3159
9.9744 | -34.4993 | -17.3257
9.9805 | -34.4998 | -17.3352
9.9866 | -34.5002 | -17.3448
9.9927 | -34.5007 | -17.3543
9.9988 | -34.5011 | -17.364
10.0049 | -34.5016 | -17.3736
10.011 | -34.502 | -17.3833
10.0171 | -34.5025 | -17.3928
10.0232 | -34.5029 | -17.4023
10.0293 | -34.5034 | -17.412
10.0354 | -34.5039 | -17.4216
10.0415 | -34.5043 | -17.4312
10.0476 | -34.5048 | -17.441
10.0537 | -34.5052 | -17.4507
10.0598 | -34.5057 | -17.4603
10.0659 | -34.5061 | -17.47
10.072 | -34.5065 | -17.4794
10.0781 | -34.507 | -17.4889
10.0842 | -34.5074 | -17.4985
10.0903 | -34.5079 | -17.5081
10.0964 | -34.5083 | -17.5174
10.1026 | -34.5088 | -17.5271
10.1087 | -34.5092 | -17.5367
10.1148 | -34.5096 | -17.5464
10.1209 | -34.5101 | -17.5558
10.127 | -34.5105 | -17.5653
10.1331 | -34.511 | -17.5748
10.1392 | -34.5114 | -17.5842
10.1453 | -34.5119 | -17.5937
10.1514 | -34.5123 | -17.6033
10.1575 | -34.5127 | -17.613
10.1636 | -34.5132 | -17.6223
10.1697 | -34.5136 | -17.632
10.1758 | -34.5141 | -17.6414
10.1819 | -34.5145 | -17.6509
10.188 | -34.5149 | -17.6605
10.1941 | -34.5154 | -17.6701
10.2002 | -34.5158 | -17.6796
10.2063 | -34.5163 | -17.6892
10.2124 | -34.5167 | -17.6988
10.2185 | -34.5172 | -17.7085
10.2246 | -34.5176 | -17.7179
10.2307 | -34.5181 | -17.7276
10.2368 | -34.5185 | -17.7368
10.2429 | -34.5189 | -17.7463
10.2491 | -34.5193 | -17.756
10.2552 | -34.5198 | -17.7656
10.2613 | -34.5202 | -17.7752
10.2674 | -34.5207 | -17.7848
10.2735 | -34.5211 | -17.7945
10.2796 | -34.5216 | -17.8041
10.2857 | -34.522 | -17.8137
10.2918 | -34.5225 | -17.8234
10.2979 | -34.5229 | -17.833
10.304 | -34.5234 | -17.8423
10.3101 | -34.5238 | -17.8519
10.3162 | -34.5243 | -17.8616
10.3223 | -34.5247 | -17.871
10.3284 | -34.5251 | -17.8804
10.3345 | -34.5256 | -17.8899
10.3406 | -34.526 | -17.8993
10.3467 | -34.5264 | -17.9088
10.3528 | -34.5269 | -17.9182
10.3589 | -34.5273 | -17.9277
10.365 | -34.5278 | -17.9373
10.3711 | -34.5282 | -17.947
10.3772 | -34.5287 | -17.9564
10.3833 | -34.5291 | -17.9659
10.3895 | -34.5296 | -17.9754
10.3956 | -34.53 | -17.985
10.4017 | -34.5305 | -17.9944
10.4078 | -34.5309 | -18.0038
10.4139 | -34.5313 | -18.0134
10.42 | -34.5318 | -18.0231
10.4261 | -34.5322 | -18.0326
10.4322 | -34.5327 | -18.0421
10.4383 | -34.5331 | -18.0516
10.4444 | -34.5335 | -18.061
10.4505 | -34.534 | -18.0705
10.4566 | -34.5344 | -18.0799
10.4627 | -34.5348 | -18.0893
10.4688 | -34.5353 | -18.0987
10.4749 | -34.5357 | -18.1082
10.481 | -34.5362 | -18.1178
10.4871 | -34.5366 | -18.1271
10.4932 | -34.5371 | -18.1365
10.4993 | -34.5375 | -18.1461
10.5054 | -34.538 | -18.1556
10.5115 | -34.5384 | -18.1651
10.5176 | -34.5389 | -18.1745
10.5237 | -34.5393 | -18.1842
10.5298 | -34.5398 | -18.1935
10.536 | -34.5402 | -18.203
10.5421 | -34.5406 | -18.2124
10.5482 | -34.5411 | -18.222
10.5543 | -34.5415 | -18.2314
10.5604 | -34.542 | -18.2407
10.5665 | -34.5424 | -18.2501
10.5726 | -34.5428 | -18.2596
10.5787 | -34.5433 | -18.269
10.5848 | -34.5437 | -18.2782
10.5909 | -34.5442 | -18.2878
10.597 | -34.5446 | -18.2973
10.6031 | -34.5451 | -18.307
10.6092 | -34.5455 | -18.3164
10.6153 | -34.5459 | -18.3261
10.6214 | -34.5464 | -18.3353
10.6275 | -34.5468 | -18.3448
10.6336 | -34.5473 | -18.3543
10.6397 | -34.5477 | -18.3638
10.6458 | -34.5481 | -18.3731
10.6519 | -34.5486 | -18.3827
10.658 | -34.549 | -18.3922
10.6641 | -34.5495 | -18.4018
10.6702 | -34.5499 | -18.4113
10.6764 | -34.5504 | -18.4208
10.6825 | -34.5508 | -18.4302
10.6886 | -34.5512 | -18.4395
10.6947 | -34.5517 | -18.4489
10.7008 | -34.5522 | -18.4583
10.7069 | -34.5526 | -18.4679
10.713 | -34.553 | -18.4773
10.7191 | -34.5535 | -18.4868
10.7252 | -34.5539 | -18.496
10.7313 | -34.5543 | -18.5054
10.7374 | -34.5548 | -18.5148
10.7435 | -34.5552 | -18.5243
10.7496 | -34.5556 | -18.5335
10.7557 | -34.5561 | -18.5431
10.7618 | -34.5565 | -18.5523
10.7679 | -34.557 | -18.5617
10.774 | -34.5574 | -18.5712
10.7801 | -34.5579 | -18.5807
10.7862 | -34.5583 | -18.5899
10.7923 | -34.5588 | -18.5995
10.7984 | -34.5592 | -18.6088
10.8045 | -34.5597 | -18.618
10.8106 | -34.5601 | -18.6275
10.8168 | -34.5605 | -18.6369
10.8229 | -34.561 | -18.6463
10.829 | -34.5614 | -18.6557
10.8351 | -34.5619 | -18.6649
10.8412 | -34.5623 | -18.6742
10.8473 | -34.5628 | -18.6835
10.8534 | -34.5632 | -18.6929
10.8595 | -34.5637 | -18.7024
10.8656 | -34.5642 | -18.7116
10.8717 | -34.5646 | -18.721
10.8778 | -34.5651 | -18.7304
10.8839 | -34.5655 | -18.7398
10.89 | -34.566 | -18.7492
10.8961 | -34.5664 | -18.7586
10.9022 | -34.5669 | -18.768
10.9083 | -34.5673 | -18.7773
10.9144 | -34.5677 | -18.7866
10.9205 | -34.5682 | -18.796
10.9266 | -34.5686 | -18.8053
10.9327 | -34.5691 | -18.8146
10.9388 | -34.5696 | -18.8238
10.9449 | -34.57 | -18.8331
10.951 | -34.5705 | -18.8425
10.9571 | -34.5709 | -18.8519
10.9633 | -34.5713 | -18.8614
10.9694 | -34.5718 | -18.871
10.9755 | -34.5723 | -18.8804
10.9816 | -34.5727 | -18.8898
10.9877 | -34.5732 | -18.899
10.9938 | -34.5736 | -18.9086
10.9999 | -34.5741 | -18.9177
11.006 | -34.5745 | -18.9269
11.0121 | -34.575 | -18.9363
11.0182 | -34.5754 | -18.9456
11.0243 | -34.5759 | -18.955
11.0304 | -34.5763 | -18.9643
11.0365 | -34.5768 | -18.9736
11.0426 | -34.5772 | -18.9829
11.0487 | -34.5777 | -18.9924
11.0548 | -34.5782 | -19.0016
11.0609 | -34.5786 | -19.011
11.067 | -34.5791 | -19.0203
11.0731 | -34.5795 | -19.0295
11.0792 | -34.58 | -19.0388
11.0853 | -34.5805 | -19.0482
11.0914 | -34.5809 | -19.0576
11.0975 | -34.5814 | -19.067
11.1037 | -34.5818 | -19.0762
11.1098 | -34.5823 | -19.0857
11.1159 | -34.5827 | -19.0948
11.122 | -34.5832 | -19.1041
11.1281 | -34.5836 | -19.1134
11.1342 | -34.5841 | -19.1228
11.1403 | -34.5845 | -19.1322
11.1464 | -34.585 | -19.1416
11.1525 | -34.5855 | -19.1507
11.1586 | -34.5859 | -19.1602
11.1647 | -34.5864 | -19.1694
11.1708 | -34.5868 | -19.1787
11.1769 | -34.5872 | -19.188
11.183 | -34.5877 | -19.1972
11.1891 | -34.5882 | -19.2065
11.1952 | -34.5887 | -19.2158
11.2013 | -34.5891 | -19.2252
11.2074 | -34.5896 | -19.2345
11.2135 | -34.59 | -19.244
11.2196 | -34.5905 | -19.2533
11.2257 | -34.591 | -19.2627
11.2318 | -34.5915 | -19.2719
11.2379 | -34.5919 | -19.2813
11.244 | -34.5924 | -19.2908
11.2502 | -34.5929 | -19.3001
11.2563 | -34.5934 | -19.3095
11.2624 | -34.5938 | -19.3187
11.2685 | -34.5943 | -19.3279
11.2746 | -34.5947 | -19.3371
11.2807 | -34.5952 | -19.3466
11.2868 | -34.5957 | -19.356
11.2929 | -34.5962 | -19.365
11.299 | -34.5966 | -19.3743
11.3051 | -34.5971 | -19.3838
11.3112 | -34.5976 | -19.3932
11.3173 | -34.598 | -19.4023
11.3234 | -34.5985 | -19.4114
11.3295 | -34.599 | -19.4208
11.3356 | -34.5994 | -19.4302
11.3417 | -34.5999 | -19.4394
11.3478 | -34.6004 | -19.4488
11.3539 | -34.6008 | -19.4582
11.36 | -34.6013 | -19.4673
11.3661 | -34.6018 | -19.4765
11.3722 | -34.6022 | -19.4855
11.3783 | -34.6027 | -19.4949
11.3844 | -34.6032 | -19.5042
11.3906 | -34.6037 | -19.5135
11.3967 | -34.6041 | -19.5227
11.4028 | -34.6046 | -19.5321
11.4089 | -34.605 | -19.5412
11.415 | -34.6055 | -19.5505
11.4211 | -34.606 | -19.5598
11.4272 | -34.6064 | -19.5691
11.4333 | -34.6069 | -19.5781
11.4394 | -34.6073 | -19.5874
11.4455 | -34.6078 | -19.5966
11.4516 | -34.6082 | -19.6059
11.4577 | -34.6087 | -19.6152
11.4638 | -34.6092 | -19.6243
11.4699 | -34.6097 | -19.6336
11.476 | -34.6101 | -19.6431
11.4821 | -34.6106 | -19.6523
11.4882 | -34.6111 | -19.6617
11.4943 | -34.6116 | -19.6711
11.5004 | -34.6121 | -19.6801
11.5065 | -34.6126 | -19.6893
11.5126 | -34.613 | -19.6985
11.5187 | -34.6135 | -19.7078
11.5248 | -34.6139 | -19.7172
11.5309 | -34.6144 | -19.7265
11.5371 | -34.6149 | -19.7356
11.5432 | -34.6154 | -19.7449
11.5493 | -34.6159 | -19.7543
11.5554 | -34.6163 | -19.7636
11.5615 | -34.6168 | -19.7729
11.5676 | -34.6173 | -19.7822
11.5737 | -34.6178 | -19.7914
11.5798 | -34.6183 | -19.8004
11.5859 | -34.6188 | -19.8096
11.592 | -34.6192 | -19.8188
11.5981 | -34.6197 | -19.828
11.6042 | -34.6202 | -19.8372
11.6103 | -34.6206 | -19.8463
11.6164 | -34.6211 | -19.8554
11.6225 | -34.6216 | -19.8646
11.6286 | -34.622 | -19.8738
11.6347 | -34.6225 | -19.8831
11.6408 | -34.623 | -19.8922
11.6469 | -34.6235 | -19.9014
11.653 | -34.6239 | -19.9106
11.6591 | -34.6244 | -19.9198
11.6652 | -34.6249 | -19.929
11.6713 | -34.6253 | -19.9381
11.6775 | -34.6258 | -19.9472
11.6836 | -34.6263 | -19.9564
11.6897 | -34.6268 | -19.9656
11.6958 | -34.6273 | -19.9749
11.7019 | -34.6277 | -19.9843
11.708 | -34.6282 | -19.9934
11.7141 | -34.6288 | -20.0024
11.7202 | -34.6292 | -20.0116
11.7263 | -34.6297 | -20.0209
11.7324 | -34.6301 | -20.0302
11.7385 | -34.6306 | -20.0395
11.7446 | -34.631 | -20.0487
11.7507 | -34.6316 | -20.0577
11.7568 | -34.632 | -20.0669
11.7629 | -34.6325 | -20.0762
11.769 | -34.633 | -20.0852
11.7751 | -34.6334 | -20.0944
11.7812 | -34.634 | -20.1034
11.7873 | -34.6344 | -20.1126
11.7934 | -34.6349 | -20.1218
11.7995 | -34.6354 | -20.131
11.8056 | -34.6359 | -20.1401
11.8117 | -34.6364 | -20.1493
11.8178 | -34.6369 | -20.1583
11.824 | -34.6373 | -20.1673
11.8301 | -34.6378 | -20.1764
11.8362 | -34.6383 | -20.1854
11.8423 | -34.6388 | -20.1946
11.8484 | -34.6393 | -20.2037
11.8545 | -34.6398 | -20.2127
11.8606 | -34.6402 | -20.2218
11.8667 | -34.6407 | -20.231
11.8728 | -34.6412 | -20.24
11.8789 | -34.6417 | -20.2491
11.885 | -34.6422 | -20.2582
11.8911 | -34.6427 | -20.2673
11.8972 | -34.6432 | -20.2763
11.9033 | -34.6436 | -20.2854
11.9094 | -34.6441 | -20.2944
11.9155 | -34.6446 | -20.3035
11.9216 | -34.6451 | -20.3125
11.9277 | -34.6456 | -20.3215
11.9338 | -34.646 | -20.3306
11.9399 | -34.6465 | -20.3398
11.946 | -34.647 | -20.3489
11.9521 | -34.6475 | -20.3579
11.9582 | -34.648 | -20.3669
11.9644 | -34.6485 | -20.3758
11.9705 | -34.649 | -20.3849
11.9766 | -34.6495 | -20.394
11.9827 | -34.65 | -20.4031
11.9888 | -34.6505 | -20.412
11.9949 | -34.651 | -20.4209
12.001 | -34.6515 | -20.43
12.0071 | -34.652 | -20.439
12.0132 | -34.6524 | -20.448
12.0193 | -34.6529 | -20.4571
12.0254 | -34.6534 | -20.4661
12.0315 | -34.6539 | -20.4749
12.0376 | -34.6544 | -20.4839
12.0437 | -34.6548 | -20.4929
12.0498 | -34.6553 | -20.502
12.0559 | -34.6558 | -20.5109
12.062 | -34.6563 | -20.5197
12.0681 | -34.6568 | -20.5288
12.0742 | -34.6573 | -20.5378
12.0803 | -34.6578 | -20.547
12.0864 | -34.6583 | -20.5561
12.0925 | -34.6588 | -20.5652
12.0986 | -34.6593 | -20.5741
12.1047 | -34.6598 | -20.5832
12.1109 | -34.6603 | -20.5921
12.117 | -34.6608 | -20.601
12.1231 | -34.6613 | -20.61
12.1292 | -34.6618 | -20.6189
12.1353 | -34.6623 | -20.6279
12.1414 | -34.6628 | -20.637
12.1475 | -34.6632 | -20.6461
12.1536 | -34.6637 | -20.6552
12.1597 | -34.6642 | -20.6642
12.1658 | -34.6647 | -20.6731
12.1719 | -34.6652 | -20.6818
12.178 | -34.6657 | -20.6908
12.1841 | -34.6662 | -20.6999
12.1902 | -34.6667 | -20.7089
12.1963 | -34.6672 | -20.718
12.2024 | -34.6677 | -20.7267
12.2085 | -34.6682 | -20.7357
12.2146 | -34.6687 | -20.7448
12.2207 | -34.6692 | -20.7537
12.2268 | -34.6697 | -20.7627
12.2329 | -34.6702 | -20.7716
12.239 | -34.6707 | -20.7802
12.2451 | -34.6712 | -20.7892
12.2513 | -34.6717 | -20.7983
12.2574 | -34.6722 | -20.8073
12.2635 | -34.6727 | -20.8162
12.2696 | -34.6732 | -20.8251
12.2757 | -34.6737 | -20.8338
12.2818 | -34.6742 | -20.8428
12.2879 | -34.6747 | -20.8518
12.294 | -34.6752 | -20.8605
12.3001 | -34.6757 | -20.8695
12.3062 | -34.6762 | -20.8785
12.3123 | -34.6768 | -20.8871
12.3184 | -34.6772 | -20.8959
12.3245 | -34.6777 | -20.9047
12.3306 | -34.6782 | -20.9136
12.3367 | -34.6787 | -20.9225
12.3428 | -34.6792 | -20.9315
12.3489 | -34.6798 | -20.9401
12.355 | -34.6802 | -20.9492
12.3611 | -34.6807 | -20.9582
12.3672 | -34.6812 | -20.967
12.3733 | -34.6817 | -20.9759
12.3794 | -34.6823 | -20.9845
12.3855 | -34.6827 | -20.9933
12.3916 | -34.6832 | -21.0022
12.3978 | -34.6838 | -21.0112
12.4039 | -34.6843 | -21.0202
12.41 | -34.6848 | -21.0291
12.4161 | -34.6853 | -21.0376
12.4222 | -34.6858 | -21.0464
12.4283 | -34.6863 | -21.0552
12.4344 | -34.6868 | -21.0642
12.4405 | -34.6873 | -21.0732
12.4466 | -34.6878 | -21.082
12.4527 | -34.6883 | -21.0907
12.4588 | -34.6888 | -21.0996
12.4649 | -34.6893 | -21.1084
12.471 | -34.6898 | -21.1174
12.4771 | -34.6902 | -21.1262
12.4832 | -34.6908 | -21.1351
12.4893 | -34.6913 | -21.1437
12.4954 | -34.6918 | -21.1525
12.5015 | -34.6923 | -21.1613
12.5076 | -34.6928 | -21.1701
12.5137 | -34.6933 | -21.179
12.5198 | -34.6938 | -21.1877
12.5259 | -34.6943 | -21.1966
12.532 | -34.6948 | -21.2055
12.5382 | -34.6953 | -21.2143
12.5443 | -34.6958 | -21.2231
12.5504 | -34.6963 | -21.232
12.5565 | -34.6968 | -21.2407
12.5626 | -34.6973 | -21.2497
12.5687 | -34.6978 | -21.2586
12.5748 | -34.6983 | -21.2674
12.5809 | -34.6988 | -21.2762
12.587 | -34.6993 | -21.2851
12.5931 | -34.6999 | -21.2937
12.5992 | -34.7004 | -21.3025
12.6053 | -34.7009 | -21.3113
12.6114 | -34.7014 | -21.3201
12.6175 | -34.7019 | -21.3291
12.6236 | -34.7024 | -21.3379
12.6297 | -34.7029 | -21.3466
12.6358 | -34.7034 | -21.3551
12.6419 | -34.7039 | -21.364
12.648 | -34.7044 | -21.373
12.6541 | -34.7049 | -21.3818
12.6602 | -34.7054 | -21.3901
12.6663 | -34.7059 | -21.3989
12.6724 | -34.7064 | -21.4075
12.6785 | -34.7069 | -21.4163
12.6847 | -34.7074 | -21.4249
12.6908 | -34.7079 | -21.4337
12.6969 | -34.7085 | -21.4424
12.703 | -34.709 | -21.4512
12.7091 | -34.7095 | -21.4599
12.7152 | -34.71 | -21.4686
12.7213 | -34.7104 | -21.4775
12.7274 | -34.711 | -21.4862
12.7335 | -34.7115 | -21.4946
12.7396 | -34.712 | -21.5032
12.7457 | -34.7125 | -21.5121
12.7518 | -34.713 | -21.5207
12.7579 | -34.7135 | -21.5296
12.764 | -34.714 | -21.5384
12.7701 | -34.7144 | -21.5472
12.7762 | -34.715 | -21.5558
12.7823 | -34.7155 | -21.5645
12.7884 | -34.716 | -21.5733
12.7945 | -34.7165 | -21.5822
12.8006 | -34.717 | -21.5908
12.8067 | -34.7175 | -21.5994
12.8128 | -34.718 | -21.608
12.8189 | -34.7185 | -21.6167
12.8251 | -34.719 | -21.6255
12.8312 | -34.7195 | -21.6343
12.8373 | -34.72 | -21.6429
12.8434 | -34.7205 | -21.6517
12.8495 | -34.721 | -21.6603
12.8556 | -34.7215 | -21.669
12.8617 | -34.722 | -21.6777
12.8678 | -34.7225 | -21.6867
12.8739 | -34.723 | -21.6952
12.88 | -34.7235 | -21.7038
12.8861 | -34.724 | -21.7124
12.8922 | -34.7245 | -21.7212
12.8983 | -34.725 | -21.7298
12.9044 | -34.7255 | -21.7383
12.9105 | -34.726 | -21.747
12.9166 | -34.7265 | -21.7557
12.9227 | -34.727 | -21.7644
12.9288 | -34.7275 | -21.7731
12.9349 | -34.728 | -21.7818
12.941 | -34.7286 | -21.7904
12.9471 | -34.7291 | -21.7992
12.9532 | -34.7296 | -21.8078
12.9593 | -34.7301 | -21.8167
12.9654 | -34.7306 | -21.8255
12.9716 | -34.7311 | -21.8343
12.9777 | -34.7316 | -21.843
12.9838 | -34.7321 | -21.8516
12.9899 | -34.7326 | -21.8603
12.996 | -34.7331 | -21.869
13.0021 | -34.7336 | -21.8777
13.0082 | -34.7342 | -21.8863
13.0143 | -34.7347 | -21.8946
13.0204 | -34.7352 | -21.9032
13.0265 | -34.7357 | -21.9118
13.0326 | -34.7362 | -21.9205
13.0387 | -34.7367 | -21.9293
13.0448 | -34.7372 | -21.938
13.0509 | -34.7377 | -21.9466
13.057 | -34.7382 | -21.9552
13.0631 | -34.7387 | -21.9638
13.0692 | -34.7393 | -21.9724
13.0753 | -34.7398 | -21.9812
13.0814 | -34.7403 | -21.9901
13.0875 | -34.7408 | -21.9985
13.0936 | -34.7413 | -22.0073
13.0997 | -34.7418 | -22.0158
13.1058 | -34.7423 | -22.0245
13.112 | -34.7428 | -22.0331
13.1181 | -34.7433 | -22.0416
13.1242 | -34.7438 | -22.0503
13.1303 | -34.7443 | -22.059
13.1364 | -34.7448 | -22.0676
13.1425 | -34.7453 | -22.0762
13.1486 | -34.7458 | -22.0848
13.1547 | -34.7463 | -22.0934
13.1608 | -34.7468 | -22.1021
13.1669 | -34.7474 | -22.1106
13.173 | -34.7479 | -22.1193
13.1791 | -34.7484 | -22.1278
13.1852 | -34.7489 | -22.1364
13.1913 | -34.7494 | -22.1449
13.1974 | -34.7499 | -22.1535
13.2035 | -34.7504 | -22.1621
13.2096 | -34.7509 | -22.1707
13.2157 | -34.7514 | -22.1793
13.2218 | -34.7519 | -22.1879
13.2279 | -34.7524 | -22.1963
13.234 | -34.7529 | -22.2049
13.2401 | -34.7535 | -22.2134
13.2462 | -34.7539 | -22.222
13.2524 | -34.7545 | -22.2306
13.2585 | -34.755 | -22.2392
13.2646 | -34.7555 | -22.2477
13.2707 | -34.756 | -22.2564
13.2768 | -34.7565 | -22.2652
13.2829 | -34.757 | -22.274
13.289 | -34.7575 | -22.2824
13.2951 | -34.758 | -22.2908
13.3012 | -34.7585 | -22.2994
13.3073 | -34.7591 | -22.308
13.3134 | -34.7595 | -22.3167
13.3195 | -34.76 | -22.3252
13.3256 | -34.7605 | -22.3337
13.3317 | -34.761 | -22.3422
13.3378 | -34.7615 | -22.351
13.3439 | -34.762 | -22.3597
13.35 | -34.7625 | -22.3684
13.3561 | -34.7631 | -22.3768
13.3622 | -34.7636 | -22.3853
13.3683 | -34.7641 | -22.3937
13.3744 | -34.7646 | -22.4022
13.3805 | -34.7651 | -22.4108
13.3866 | -34.7656 | -22.4193
13.3927 | -34.766 | -22.428
13.3989 | -34.7666 | -22.4365
13.405 | -34.7671 | -22.4451
13.4111 | -34.7676 | -22.4536
13.4172 | -34.7681 | -22.4622
13.4233 | -34.7686 | -22.4708
13.4294 | -34.7691 | -22.4796
13.4355 | -34.7696 | -22.488
13.4416 | -34.7701 | -22.4965
13.4477 | -34.7706 | -22.5051
13.4538 | -34.7711 | -22.5136
13.4599 | -34.7716 | -22.5223
13.466 | -34.7721 | -22.5309
13.4721 | -34.7726 | -22.5393
13.4782 | -34.7731 | -22.5478
13.4843 | -34.7736 | -22.5565
13.4904 | -34.7741 | -22.565
13.4965 | -34.7746 | -22.5736
13.5026 | -34.7751 | -22.5822
13.5087 | -34.7756 | -22.5906
13.5148 | -34.7761 | -22.5994
13.5209 | -34.7766 | -22.6081
13.527 | -34.7771 | -22.6166
13.5331 | -34.7777 | -22.6251
13.5393 | -34.7782 | -22.6335
13.5454 | -34.7787 | -22.642
13.5515 | -34.7792 | -22.6507
13.5576 | -34.7797 | -22.6593
13.5637 | -34.7802 | -22.6679
13.5698 | -34.7807 | -22.6764
13.5759 | -34.7812 | -22.6845
13.582 | -34.7817 | -22.693
13.5881 | -34.7822 | -22.7015
13.5942 | -34.7827 | -22.7102
13.6003 | -34.7833 | -22.7188
13.6064 | -34.7838 | -22.7275
13.6125 | -34.7843 | -22.7362
13.6186 | -34.7848 | -22.7446
13.6247 | -34.7853 | -22.7532
13.6308 | -34.7858 | -22.7618
13.6369 | -34.7863 | -22.7704
13.643 | -34.7868 | -22.7791
13.6491 | -34.7873 | -22.7876
13.6552 | -34.7879 | -22.7961
13.6613 | -34.7884 | -22.8047
13.6674 | -34.7888 | -22.8132
13.6735 | -34.7893 | -22.8218
13.6796 | -34.7899 | -22.8303
13.6858 | -34.7903 | -22.8388
13.6919 | -34.7909 | -22.8474
13.698 | -34.7914 | -22.8559
13.7041 | -34.7919 | -22.8646
13.7102 | -34.7924 | -22.8731
13.7163 | -34.7929 | -22.8815
13.7224 | -34.7934 | -22.8901
13.7285 | -34.7939 | -22.8986
13.7346 | -34.7944 | -22.9072
13.7407 | -34.795 | -22.9157
13.7468 | -34.7955 | -22.9243
13.7529 | -34.796 | -22.9327
13.759 | -34.7965 | -22.9413
13.7651 | -34.797 | -22.9498
13.7712 | -34.7975 | -22.9584
13.7773 | -34.798 | -22.9669
13.7834 | -34.7985 | -22.9755
13.7895 | -34.799 | -22.9839
13.7956 | -34.7995 | -22.9925
13.8017 | -34.8 | -23.001
13.8078 | -34.8005 | -23.0094
13.8139 | -34.801 | -23.0179
13.82 | -34.8015 | -23.0263
13.8262 | -34.802 | -23.0348
13.8323 | -34.8025 | -23.0434
13.8384 | -34.803 | -23.052
13.8445 | -34.8036 | -23.0605
13.8506 | -34.8041 | -23.0691
13.8567 | -34.8046 | -23.0776
13.8628 | -34.8051 | -23.0862
13.8689 | -34.8056 | -23.0948
13.875 | -34.8061 | -23.1032
13.8811 | -34.8066 | -23.1118
13.8872 | -34.807 | -23.1203
13.8933 | -34.8076 | -23.1286
13.8994 | -34.808 | -23.1372
13.9055 | -34.8086 | -23.1459
13.9116 | -34.8091 | -23.1544
13.9177 | -34.8096 | -23.1629
13.9238 | -34.8101 | -23.1714
13.9299 | -34.8106 | -23.1799
13.936 | -34.8111 | -23.1885
13.9421 | -34.8115 | -23.197
13.9482 | -34.812 | -23.2055
13.9543 | -34.8125 | -23.2141
13.9604 | -34.813 | -23.2226
13.9665 | -34.8136 | -23.2311
13.9727 | -34.8141 | -23.2398
13.9788 | -34.8146 | -23.2482
13.9849 | -34.8151 | -23.2568
13.991 | -34.8156 | -23.2653
13.9971 | -34.8161 | -23.2737
14.0032 | -34.8166 | -23.2823
14.0093 | -34.8171 | -23.2908
14.0154 | -34.8176 | -23.2991
14.0215 | -34.8181 | -23.3075
14.0276 | -34.8186 | -23.3161
14.0337 | -34.8192 | -23.3246
14.0398 | -34.8197 | -23.333
14.0459 | -34.8202 | -23.3415
14.052 | -34.8207 | -23.3501
14.0581 | -34.8212 | -23.3587
14.0642 | -34.8217 | -23.3672
14.0703 | -34.8222 | -23.3756
14.0764 | -34.8227 | -23.384
14.0825 | -34.8232 | -23.3926
14.0886 | -34.8237 | -23.4011
14.0947 | -34.8242 | -23.4096
14.1008 | -34.8247 | -23.4181
14.1069 | -34.8252 | -23.4265
14.1131 | -34.8257 | -23.4351
14.1192 | -34.8262 | -23.4435
14.1253 | -34.8267 | -23.4518
14.1314 | -34.8272 | -23.4603
14.1375 | -34.8278 | -23.4688
14.1436 | -34.8283 | -23.4772
14.1497 | -34.8288 | -23.4856
14.1558 | -34.8293 | -23.4941
14.1619 | -34.8298 | -23.5026
14.168 | -34.8303 | -23.511
14.1741 | -34.8309 | -23.5195
14.1802 | -34.8314 | -23.5279
14.1863 | -34.8319 | -23.5364
14.1924 | -34.8324 | -23.5449
14.1985 | -34.8329 | -23.5533
14.2046 | -34.8334 | -23.5619
14.2107 | -34.8339 | -23.5702
14.2168 | -34.8344 | -23.5787
14.2229 | -34.8349 | -23.5871
14.229 | -34.8354 | -23.5956
14.2351 | -34.8359 | -23.604
14.2412 | -34.8364 | -23.6126
14.2473 | -34.8369 | -23.6209
14.2534 | -34.8374 | -23.6292
14.2596 | -34.8379 | -23.6375
14.2657 | -34.8384 | -23.646
14.2718 | -34.8389 | -23.6545
14.2779 | -34.8394 | -23.663
14.284 | -34.84 | -23.6716
14.2901 | -34.8405 | -23.68
14.2962 | -34.841 | -23.6886
14.3023 | -34.8415 | -23.697
14.3084 | -34.842 | -23.7055
14.3145 | -34.8425 | -23.714
14.3206 | -34.843 | -23.7226
14.3267 | -34.8435 | -23.7311
14.3328 | -34.8441 | -23.7395
14.3389 | -34.8446 | -23.748
14.345 | -34.8451 | -23.7565
14.3511 | -34.8456 | -23.7649
14.3572 | -34.8461 | -23.7735
14.3633 | -34.8466 | -23.7821
14.3694 | -34.8471 | -23.7905
14.3755 | -34.8476 | -23.7989
14.3816 | -34.8482 | -23.8075
14.3877 | -34.8487 | -23.8159
14.3938 | -34.8492 | -23.8245
14.4 | -34.8497 | -23.8329
14.4061 | -34.8502 | -23.8413
14.4122 | -34.8507 | -23.8498
14.4183 | -34.8512 | -23.8582
14.4244 | -34.8517 | -23.8668
14.4305 | -34.8522 | -23.8751
14.4366 | -34.8528 | -23.8835
14.4427 | -34.8533 | -23.8921
14.4488 | -34.8538 | -23.9006
14.4549 | -34.8543 | -23.909
14.461 | -34.8549 | -23.9173
14.4671 | -34.8554 | -23.9259
14.4732 | -34.8559 | -23.9345
14.4793 | -34.8564 | -23.943
14.4854 | -34.8569 | -23.9514
14.4915 | -34.8574 | -23.9595
14.4976 | -34.858 | -23.9679
14.5037 | -34.8585 | -23.9762
14.5098 | -34.859 | -23.9848
14.5159 | -34.8596 | -23.9933
14.522 | -34.8601 | -24.0016
14.5281 | -34.8606 | -24.0099
14.5342 | -34.8611 | -24.0184
14.5403 | -34.8616 | -24.0268
14.5465 | -34.8621 | -24.0354
14.5526 | -34.8626 | -24.044
14.5587 | -34.8632 | -24.0525
14.5648 | -34.8637 | -24.0607
14.5709 | -34.8642 | -24.0691
14.577 | -34.8647 | -24.0775
14.5831 | -34.8652 | -24.086
14.5892 | -34.8658 | -24.0944
14.5953 | -34.8663 | -24.1026
14.6014 | -34.8668 | -24.1111
14.6075 | -34.8673 | -24.1195
14.6136 | -34.8678 | -24.128
14.6197 | -34.8683 | -24.1366
14.6258 | -34.8688 | -24.145
14.6319 | -34.8693 | -24.1534
14.638 | -34.8698 | -24.1618
14.6441 | -34.8703 | -24.1705
14.6502 | -34.8708 | -24.1791
14.6563 | -34.8714 | -24.1875
14.6624 | -34.8719 | -24.196
14.6685 | -34.8725 | -24.2042
14.6746 | -34.873 | -24.2126
14.6807 | -34.8735 | -24.2209
14.6869 | -34.874 | -24.2292
14.693 | -34.8746 | -24.2375
14.6991 | -34.8751 | -24.2458
14.7052 | -34.8756 | -24.2542
14.7113 | -34.8761 | -24.2627
14.7174 | -34.8766 | -24.2712
14.7235 | -34.8772 | -24.2795
14.7296 | -34.8777 | -24.2879
14.7357 | -34.8782 | -24.2962
14.7418 | -34.8787 | -24.3045
14.7479 | -34.8793 | -24.3129
14.754 | -34.8798 | -24.3213
14.7601 | -34.8803 | -24.3297
14.7662 | -34.8808 | -24.338
14.7723 | -34.8813 | -24.3463
14.7784 | -34.8818 | -24.3548
14.7845 | -34.8824 | -24.3632
14.7906 | -34.8829 | -24.3717
14.7967 | -34.8834 | -24.3801
14.8028 | -34.8839 | -24.3885
14.8089 | -34.8845 | -24.3968
14.815 | -34.885 | -24.4052
14.8211 | -34.8855 | -24.4135
14.8272 | -34.886 | -24.422
14.8334 | -34.8866 | -24.4304
14.8395 | -34.8871 | -24.4387
14.8456 | -34.8876 | -24.4471
14.8517 | -34.8881 | -24.4556
14.8578 | -34.8886 | -24.464
14.8639 | -34.8892 | -24.4726
14.87 | -34.8897 | -24.481
14.8761 | -34.8903 | -24.4894
14.8822 | -34.8908 | -24.4979
14.8883 | -34.8913 | -24.5062
14.8944 | -34.8918 | -24.5145
14.9005 | -34.8923 | -24.5229
14.9066 | -34.8928 | -24.5313
14.9127 | -34.8934 | -24.5396
14.9188 | -34.8939 | -24.5479
14.9249 | -34.8944 | -24.5563
14.931 | -34.8949 | -24.5648
14.9371 | -34.8955 | -24.5731
14.9432 | -34.896 | -24.5815
14.9493 | -34.8965 | -24.5898
14.9554 | -34.8971 | -24.5982
14.9615 | -34.8976 | -24.6066
14.9676 | -34.8981 | -24.6152
14.9738 | -34.8986 | -24.6235
14.9799 | -34.8991 | -24.632
14.986 | -34.8996 | -24.6403
14.9921 | -34.9001 | -24.6486
14.9982 | -34.9006 | -24.6569
15.0043 | -34.9012 | -24.6652
15.0104 | -34.9017 | -24.6737
15.0165 | -34.9022 | -24.6823
15.0226 | -34.9028 | -24.6906
15.0287 | -34.9033 | -24.6991
15.0348 | -34.9038 | -24.7076
15.0409 | -34.9043 | -24.7159
15.047 | -34.9049 | -24.7244
15.0531 | -34.9054 | -24.7327
15.0592 | -34.9059 | -24.7411
15.0653 | -34.9065 | -24.7494
15.0714 | -34.907 | -24.7576
15.0775 | -34.9075 | -24.7659
15.0836 | -34.908 | -24.7744
15.0897 | -34.9086 | -24.7827
15.0958 | -34.9091 | -24.791
15.1019 | -34.9097 | -24.7995
15.108 | -34.9102 | -24.8078
15.1141 | -34.9107 | -24.8162
15.1203 | -34.9112 | -24.8246
15.1264 | -34.9118 | -24.833
15.1325 | -34.9123 | -24.8414
15.1386 | -34.9128 | -24.8496
15.1447 | -34.9134 | -24.858
15.1508 | -34.9139 | -24.8663
15.1569 | -34.9144 | -24.8747
15.163 | -34.915 | -24.8828
15.1691 | -34.9155 | -24.8911
15.1752 | -34.9161 | -24.8995
15.1813 | -34.9166 | -24.9077
15.1874 | -34.9171 | -24.9161
15.1935 | -34.9177 | -24.9242
15.1996 | -34.9182 | -24.9325
15.2057 | -34.9187 | -24.9407
15.2118 | -34.9193 | -24.9491
15.2179 | -34.9198 | -24.9576
15.224 | -34.9203 | -24.9659
15.2301 | -34.9209 | -24.9741
15.2362 | -34.9214 | -24.9825
15.2423 | -34.9219 | -24.9908
15.2484 | -34.9225 | -24.9992
15.2545 | -34.923 | -25.0076
15.2607 | -34.9236 | -25.0159
15.2668 | -34.9241 | -25.0242
15.2729 | -34.9247 | -25.0324
15.279 | -34.9252 | -25.041
15.2851 | -34.9258 | -25.0493
15.2912 | -34.9263 | -25.0576
15.2973 | -34.9268 | -25.0661
15.3034 | -34.9273 | -25.0742
15.3095 | -34.9279 | -25.0826
15.3156 | -34.9284 | -25.0909
15.3217 | -34.929 | -25.0993
15.3278 | -34.9295 | -25.1077
15.3339 | -34.93 | -25.1159
15.34 | -34.9306 | -25.1241
15.3461 | -34.9311 | -25.1324
15.3522 | -34.9316 | -25.1406
15.3583 | -34.9322 | -25.1491
15.3644 | -34.9327 | -25.1573
15.3705 | -34.9333 | -25.1654
15.3766 | -34.9338 | -25.1738
15.3827 | -34.9343 | -25.182
15.3888 | -34.9349 | -25.1903
15.3949 | -34.9354 | -25.1985
15.401 | -34.9359 | -25.2068
15.4072 | -34.9364 | -25.2148
15.4133 | -34.937 | -25.2231
15.4194 | -34.9375 | -25.2314
15.4255 | -34.938 | -25.2396
15.4316 | -34.9386 | -25.248
15.4377 | -34.9391 | -25.2566
15.4438 | -34.9396 | -25.2647
15.4499 | -34.9402 | -25.2729
15.456 | -34.9407 | -25.2811
15.4621 | -34.9413 | -25.2893
15.4682 | -34.9418 | -25.2977
15.4743 | -34.9423 | -25.3058
15.4804 | -34.9429 | -25.3141
15.4865 | -34.9434 | -25.3223
15.4926 | -34.9439 | -25.3306
15.4987 | -34.9445 | -25.3387
15.5048 | -34.945 | -25.347
15.5109 | -34.9455 | -25.3551
15.517 | -34.9461 | -25.3634
15.5231 | -34.9466 | -25.3717
15.5292 | -34.9471 | -25.3798
15.5353 | -34.9477 | -25.3882
15.5414 | -34.9482 | -25.3965
15.5476 | -34.9488 | -25.4044
15.5537 | -34.9493 | -25.4127
15.5598 | -34.9498 | -25.4211
15.5659 | -34.9504 | -25.4293
15.572 | -34.9509 | -25.4375
15.5781 | -34.9515 | -25.4457
15.5842 | -34.952 | -25.4537
15.5903 | -34.9526 | -25.4619
15.5964 | -34.9531 | -25.4701
15.6025 | -34.9536 | -25.4783
15.6086 | -34.9542 | -25.4866
15.6147 | -34.9547 | -25.495
15.6208 | -34.9553 | -25.5031
15.6269 | -34.9558 | -25.5114
15.633 | -34.9563 | -25.5199
15.6391 | -34.9569 | -25.5281
15.6452 | -34.9574 | -25.5364
15.6513 | -34.958 | -25.5442
15.6574 | -34.9585 | -25.5526
15.6635 | -34.9591 | -25.5607
15.6696 | -34.9596 | -25.569
15.6757 | -34.9602 | -25.5773
15.6818 | -34.9607 | -25.5855
15.688 | -34.9612 | -25.5937
15.6941 | -34.9618 | -25.6018
15.7002 | -34.9623 | -25.61
15.7063 | -34.9628 | -25.6182
15.7124 | -34.9634 | -25.6265
15.7185 | -34.9639 | -25.6345
15.7246 | -34.9644 | -25.6426
15.7307 | -34.965 | -25.6508
15.7368 | -34.9655 | -25.659
15.7429 | -34.9661 | -25.6672
15.749 | -34.9666 | -25.6754
15.7551 | -34.9672 | -25.6838
15.7612 | -34.9678 | -25.6918
15.7673 | -34.9683 | -25.7
15.7734 | -34.9689 | -25.708
15.7795 | -34.9694 | -25.7162
15.7856 | -34.9699 | -25.7242
15.7917 | -34.9705 | -25.7321
15.7978 | -34.971 | -25.7404
15.8039 | -34.9716 | -25.7486
15.81 | -34.9721 | -25.7568
15.8161 | -34.9726 | -25.7649
15.8222 | -34.9732 | -25.773
15.8283 | -34.9738 | -25.7809
15.8345 | -34.9743 | -25.7891
15.8406 | -34.9748 | -25.7972
15.8467 | -34.9754 | -25.8054
15.8528 | -34.9759 | -25.8137
15.8589 | -34.9765 | -25.8219
15.865 | -34.977 | -25.8299
15.8711 | -34.9775 | -25.8381
15.8772 | -34.9781 | -25.8461
15.8833 | -34.9786 | -25.8544
15.8894 | -34.9792 | -25.8627
15.8955 | -34.9797 | -25.8709
15.9016 | -34.9803 | -25.8788
15.9077 | -34.9808 | -25.8868
15.9138 | -34.9813 | -25.8949
15.9199 | -34.9819 | -25.903
15.926 | -34.9824 | -25.9111
15.9321 | -34.9829 | -25.9194
15.9382 | -34.9835 | -25.9275
15.9443 | -34.984 | -25.9355
15.9504 | -34.9846 | -25.9439
15.9565 | -34.9851 | -25.952
15.9626 | -34.9856 | -25.9601
15.9687 | -34.9861 | -25.9682
15.9749 | -34.9867 | -25.976
15.981 | -34.9872 | -25.9842
15.9871 | -34.9877 | -25.9923
15.9932 | -34.9882 | -26.0002
15.9993 | -34.9888 | -26.0082
16.0054 | -34.9894 | -26.0159
16.0115 | -34.9899 | -26.0241
16.0176 | -34.9905 | -26.0324
16.0237 | -34.991 | -26.0406
16.0298 | -34.9915 | -26.0485
16.0359 | -34.9921 | -26.0565
16.042 | -34.9926 | -26.0645
16.0481 | -34.9931 | -26.0726
16.0542 | -34.9937 | -26.0809
16.0603 | -34.9942 | -26.0889
16.0664 | -34.9948 | -26.097
16.0725 | -34.9953 | -26.1048
16.0786 | -34.9959 | -26.1128
16.0847 | -34.9964 | -26.121
16.0908 | -34.9969 | -26.129
16.0969 | -34.9975 | -26.1371
16.103 | -34.998 | -26.1452
16.1091 | -34.9986 | -26.1532
16.1152 | -34.9991 | -26.1614
16.1214 | -34.9997 | -26.1696
16.1275 | -35.0002 | -26.1776
16.1336 | -35.0008 | -26.1859
16.1397 | -35.0013 | -26.194
16.1458 | -35.0019 | -26.2018
16.1519 | -35.0024 | -26.2101
16.158 | -35.0029 | -26.2182
16.1641 | -35.0035 | -26.2263
16.1702 | -35.004 | -26.2342
16.1763 | -35.0045 | -26.2424
16.1824 | -35.0051 | -26.2503
16.1885 | -35.0056 | -26.2583
16.1946 | -35.0061 | -26.2662
16.2007 | -35.0067 | -26.2741
16.2068 | -35.0072 | -26.282
16.2129 | -35.0078 | -26.2902
16.219 | -35.0083 | -26.2978
16.2251 | -35.0089 | -26.3056
16.2312 | -35.0094 | -26.3136
16.2373 | -35.01 | -26.3215
16.2434 | -35.0105 | -26.3294
16.2495 | -35.0111 | -26.3372
16.2556 | -35.0116 | -26.3454
16.2618 | -35.0122 | -26.3534
16.2679 | -35.0127 | -26.3615
16.274 | -35.0133 | -26.3695
16.2801 | -35.0138 | -26.3773
16.2862 | -35.0143 | -26.3851
16.2923 | -35.0149 | -26.3931
16.2984 | -35.0154 | -26.4012
16.3045 | -35.0159 | -26.4091
16.3106 | -35.0165 | -26.4171
16.3167 | -35.017 | -26.4251
16.3228 | -35.0176 | -26.4329
16.3289 | -35.0181 | -26.441
16.335 | -35.0186 | -26.4489
16.3411 | -35.0192 | -26.4568
16.3472 | -35.0197 | -26.4649
16.3533 | -35.0203 | -26.4728
16.3594 | -35.0208 | -26.4807
16.3655 | -35.0213 | -26.4886
16.3716 | -35.0218 | -26.4963
16.3777 | -35.0224 | -26.5043
16.3838 | -35.0229 | -26.5123
16.3899 | -35.0234 | -26.5201
16.396 | -35.024 | -26.5281
16.4021 | -35.0246 | -26.536
16.4083 | -35.0251 | -26.5439
16.4144 | -35.0256 | -26.5521
16.4205 | -35.0262 | -26.5602
16.4266 | -35.0268 | -26.5679
16.4327 | -35.0273 | -26.5758
16.4388 | -35.0278 | -26.5837
16.4449 | -35.0284 | -26.5916
16.451 | -35.0289 | -26.5997
16.4571 | -35.0295 | -26.6078
16.4632 | -35.03 | -26.6156
16.4693 | -35.0305 | -26.6237
16.4754 | -35.0311 | -26.6318
16.4815 | -35.0316 | -26.6399
16.4876 | -35.0321 | -26.6476
16.4937 | -35.0326 | -26.6554
16.4998 | -35.0332 | -26.6632
16.5059 | -35.0338 | -26.6711
16.512 | -35.0343 | -26.6792
16.5181 | -35.0348 | -26.6874
16.5242 | -35.0354 | -26.6954
16.5303 | -35.0359 | -26.7033
16.5364 | -35.0364 | -26.7112
16.5425 | -35.037 | -26.719
16.5487 | -35.0375 | -26.7268
16.5548 | -35.038 | -26.7348
16.5609 | -35.0386 | -26.7428
16.567 | -35.0392 | -26.7505
16.5731 | -35.0397 | -26.7586
16.5792 | -35.0403 | -26.7665
16.5853 | -35.0408 | -26.7744
16.5914 | -35.0414 | -26.7823
16.5975 | -35.0419 | -26.7903
16.6036 | -35.0424 | -26.798
16.6097 | -35.0429 | -26.8059
16.6158 | -35.0434 | -26.8138
16.6219 | -35.044 | -26.8218
16.628 | -35.0445 | -26.8298
16.6341 | -35.045 | -26.8376
16.6402 | -35.0456 | -26.8453
16.6463 | -35.0461 | -26.8533
16.6524 | -35.0466 | -26.8611
16.6585 | -35.0472 | -26.8691
16.6646 | -35.0477 | -26.8769
16.6707 | -35.0482 | -26.8848
16.6768 | -35.0488 | -26.8927
16.6829 | -35.0493 | -26.9005
16.689 | -35.0498 | -26.9084
16.6952 | -35.0503 | -26.9163
16.7013 | -35.0509 | -26.9241
16.7074 | -35.0514 | -26.9318
16.7135 | -35.052 | -26.9396
16.7196 | -35.0525 | -26.9473
16.7257 | -35.0531 | -26.9552
16.7318 | -35.0536 | -26.963
16.7379 | -35.0541 | -26.9709
16.744 | -35.0547 | -26.9788
16.7501 | -35.0552 | -26.9867
16.7562 | -35.0557 | -26.9947
16.7623 | -35.0562 | -27.0023
16.7684 | -35.0567 | -27.0103
16.7745 | -35.0573 | -27.0184
16.7806 | -35.0579 | -27.0259
16.7867 | -35.0584 | -27.0336
16.7928 | -35.059 | -27.0417
16.7989 | -35.0595 | -27.0496
16.805 | -35.06 | -27.0575
16.8111 | -35.0605 | -27.0653
16.8172 | -35.0611 | -27.0732
16.8233 | -35.0617 | -27.0811
16.8294 | -35.0622 | -27.089
16.8356 | -35.0627 | -27.0969
16.8417 | -35.0633 | -27.1048
16.8478 | -35.0638 | -27.1126
16.8539 | -35.0643 | -27.1203
16.86 | -35.0649 | -27.1281
16.8661 | -35.0654 | -27.1359
16.8722 | -35.066 | -27.1437
16.8783 | -35.0665 | -27.1516
16.8844 | -35.0671 | -27.1592
16.8905 | -35.0676 | -27.1671
16.8966 | -35.0681 | -27.1747
16.9027 | -35.0687 | -27.1825
16.9088 | -35.0692 | -27.1906
16.9149 | -35.0698 | -27.1984
16.921 | -35.0703 | -27.206
16.9271 | -35.0708 | -27.2138
16.9332 | -35.0714 | -27.2216
16.9393 | -35.0719 | -27.2294
16.9454 | -35.0724 | -27.2371
16.9515 | -35.0729 | -27.2449
16.9576 | -35.0734 | -27.2524
16.9637 | -35.074 | -27.2601
16.9698 | -35.0745 | -27.2679
16.9759 | -35.0751 | -27.2755
16.9821 | -35.0756 | -27.2836
16.9882 | -35.0762 | -27.291
16.9943 | -35.0767 | -27.2985
17.0004 | -35.0772 | -27.3063
17.0065 | -35.0778 | -27.3141
17.0126 | -35.0783 | -27.3218
17.0187 | -35.0788 | -27.3297
17.0248 | -35.0793 | -27.3376
17.0309 | -35.0799 | -27.3455
17.037 | -35.0804 | -27.3532
17.0431 | -35.081 | -27.3609
17.0492 | -35.0815 | -27.3686
17.0553 | -35.082 | -27.3763
17.0614 | -35.0825 | -27.3839
17.0675 | -35.0831 | -27.3917
17.0736 | -35.0836 | -27.3996
17.0797 | -35.0842 | -27.4073
17.0858 | -35.0847 | -27.4153
17.0919 | -35.0853 | -27.4233
17.098 | -35.0859 | -27.4307
17.1041 | -35.0864 | -27.4385
17.1102 | -35.087 | -27.4463
17.1163 | -35.0875 | -27.4542
17.1225 | -35.0881 | -27.4618
17.1286 | -35.0886 | -27.4694
17.1347 | -35.0892 | -27.4772
17.1408 | -35.0897 | -27.485
17.1469 | -35.0902 | -27.4927
17.153 | -35.0908 | -27.5003
17.1591 | -35.0913 | -27.508
17.1652 | -35.0919 | -27.5155
17.1713 | -35.0924 | -27.5232
17.1774 | -35.093 | -27.5308
17.1835 | -35.0935 | -27.5384
17.1896 | -35.094 | -27.546
17.1957 | -35.0946 | -27.5536
17.2018 | -35.0951 | -27.5607
17.2079 | -35.0957 | -27.5687
17.214 | -35.0962 | -27.5766
17.2201 | -35.0968 | -27.5843
17.2262 | -35.0973 | -27.5921
17.2323 | -35.0978 | -27.5998
17.2384 | -35.0984 | -27.6078
17.2445 | -35.099 | -27.6153
17.2506 | -35.0995 | -27.6229
17.2567 | -35.1 | -27.6307
17.2628 | -35.1006 | -27.6385
17.269 | -35.1011 | -27.646
17.2751 | -35.1016 | -27.6537
17.2812 | -35.1021 | -27.6615
17.2873 | -35.1027 | -27.6694
17.2934 | -35.1032 | -27.6771
17.2995 | -35.1038 | -27.6846
17.3056 | -35.1043 | -27.6921
17.3117 | -35.1049 | -27.6999
17.3178 | -35.1054 | -27.7077
17.3239 | -35.1059 | -27.7153
17.33 | -35.1065 | -27.723
17.3361 | -35.107 | -27.7305
17.3422 | -35.1076 | -27.7382
17.3483 | -35.1081 | -27.746
17.3544 | -35.1086 | -27.7535
17.3605 | -35.1091 | -27.7612
17.3666 | -35.1096 | -27.7688
17.3727 | -35.1101 | -27.7766
17.3788 | -35.1107 | -27.7839
17.3849 | -35.1112 | -27.7915
17.391 | -35.1117 | -27.7991
17.3971 | -35.1122 | -27.8066
17.4032 | -35.1128 | -27.8142
17.4094 | -35.1133 | -27.8218
17.4155 | -35.1139 | -27.8294
17.4216 | -35.1144 | -27.837
17.4277 | -35.115 | -27.8448
17.4338 | -35.1155 | -27.8526
17.4399 | -35.116 | -27.8602
17.446 | -35.1165 | -27.8676
17.4521 | -35.1171 | -27.8752
17.4582 | -35.1176 | -27.8826
17.4643 | -35.1181 | -27.8901
17.4704 | -35.1187 | -27.8979
17.4765 | -35.1192 | -27.9056
17.4826 | -35.1198 | -27.9132
17.4887 | -35.1203 | -27.9207
17.4948 | -35.1208 | -27.9283
17.5009 | -35.1214 | -27.936
17.507 | -35.122 | -27.9437
17.5131 | -35.1225 | -27.9515
17.5192 | -35.123 | -27.9591
17.5253 | -35.1235 | -27.9668
17.5314 | -35.124 | -27.9747
17.5375 | -35.1245 | -27.9822
17.5436 | -35.125 | -27.9898
17.5497 | -35.1256 | -27.9972
17.5559 | -35.1261 | -28.0046
17.562 | -35.1267 | -28.0122
17.5681 | -35.1272 | -28.0196
17.5742 | -35.1277 | -28.0274
17.5803 | -35.1283 | -28.035
17.5864 | -35.1288 | -28.0427
17.5925 | -35.1293 | -28.0504
17.5986 | -35.1299 | -28.0576
17.6047 | -35.1304 | -28.0653
17.6108 | -35.1309 | -28.0729
17.6169 | -35.1315 | -28.0803
17.623 | -35.132 | -28.0877
17.6291 | -35.1326 | -28.0952
17.6352 | -35.1331 | -28.1027
17.6413 | -35.1337 | -28.1104
17.6474 | -35.1342 | -28.118
17.6535 | -35.1348 | -28.1255
17.6596 | -35.1353 | -28.1327
17.6657 | -35.1359 | -28.1403
17.6718 | -35.1364 | -28.148
17.6779 | -35.1369 | -28.1555
17.684 | -35.1375 | -28.1632
17.6901 | -35.1381 | -28.1704
17.6963 | -35.1385 | -28.1777
17.7024 | -35.1391 | -28.1853
17.7085 | -35.1396 | -28.1927
17.7146 | -35.1401 | -28.2003
17.7207 | -35.1406 | -28.2078
17.7268 | -35.1412 | -28.215
17.7329 | -35.1417 | -28.2225
17.739 | -35.1422 | -28.2302
17.7451 | -35.1428 | -28.2379
17.7512 | -35.1433 | -28.2451
17.7573 | -35.1438 | -28.2527
17.7634 | -35.1443 | -28.26
17.7695 | -35.1449 | -28.2675
17.7756 | -35.1454 | -28.2752
17.7817 | -35.1459 | -28.2829
17.7878 | -35.1464 | -28.2906
17.7939 | -35.147 | -28.2981
17.8 | -35.1475 | -28.3057
17.8061 | -35.148 | -28.3129
17.8122 | -35.1485 | -28.3205
17.8183 | -35.1491 | -28.328
17.8244 | -35.1496 | -28.3354
17.8305 | -35.1502 | -28.3428
17.8366 | -35.1507 | -28.35
17.8428 | -35.1512 | -28.3574
17.8489 | -35.1518 | -28.3649
17.855 | -35.1523 | -28.3722
17.8611 | -35.1528 | -28.3796
17.8672 | -35.1533 | -28.3869
17.8733 | -35.1539 | -28.3946
17.8794 | -35.1544 | -28.4019
17.8855 | -35.1549 | -28.4093
17.8916 | -35.1555 | -28.4169
17.8977 | -35.156 | -28.4247
17.9038 | -35.1566 | -28.4318
17.9099 | -35.1571 | -28.4394
17.916 | -35.1576 | -28.4469
17.9221 | -35.1582 | -28.4545
17.9282 | -35.1587 | -28.4619
17.9343 | -35.1592 | -28.4695
17.9404 | -35.1598 | -28.4768
17.9465 | -35.1603 | -28.4843
17.9526 | -35.1608 | -28.492
17.9587 | -35.1614 | -28.4995
17.9648 | -35.1619 | -28.5068
17.9709 | -35.1625 | -28.5142
17.977 | -35.163 | -28.5216
17.9832 | -35.1635 | -28.5291
17.9893 | -35.164 | -28.5366
17.9954 | -35.1645 | -28.5442
18.0015 | -35.165 | -28.5517
18.0076 | -35.1656 | -28.559
18.0137 | -35.1661 | -28.5664
18.0198 | -35.1666 | -28.5739
18.0259 | -35.1672 | -28.5813
18.032 | -35.1677 | -28.5886
18.0381 | -35.1682 | -28.5959
18.0442 | -35.1687 | -28.6033
18.0503 | -35.1692 | -28.6105
18.0564 | -35.1698 | -28.6179
18.0625 | -35.1703 | -28.6253
18.0686 | -35.1709 | -28.6326
18.0747 | -35.1714 | -28.6404
18.0808 | -35.1719 | -28.6478
18.0869 | -35.1724 | -28.6553
18.093 | -35.173 | -28.6628
18.0991 | -35.1734 | -28.6702
18.1052 | -35.1739 | -28.678
18.1113 | -35.1745 | -28.6854
18.1174 | -35.175 | -28.6925
18.1236 | -35.1756 | -28.7
18.1297 | -35.1761 | -28.7075
18.1358 | -35.1767 | -28.7152
18.1419 | -35.1772 | -28.7225
18.148 | -35.1778 | -28.73
18.1541 | -35.1783 | -28.7372
18.1602 | -35.1789 | -28.7447
18.1663 | -35.1794 | -28.7523
18.1724 | -35.18 | -28.7601
18.1785 | -35.1805 | -28.7677
18.1846 | -35.181 | -28.7749
18.1907 | -35.1816 | -28.7824
18.1968 | -35.1821 | -28.7897
18.2029 | -35.1826 | -28.7972
18.209 | -35.1832 | -28.8044
18.2151 | -35.1837 | -28.8117
18.2212 | -35.1843 | -28.8189
18.2273 | -35.1848 | -28.8262
18.2334 | -35.1853 | -28.8335
18.2395 | -35.1858 | -28.8409
18.2456 | -35.1864 | -28.8486
18.2517 | -35.1869 | -28.8561
18.2578 | -35.1874 | -28.8631
18.2639 | -35.188 | -28.8706
18.2701 | -35.1885 | -28.8779
18.2762 | -35.189 | -28.8848
18.2823 | -35.1895 | -28.8919
18.2884 | -35.1901 | -28.8991
18.2945 | -35.1907 | -28.9059
18.3006 | -35.1912 | -28.9132
18.3067 | -35.1918 | -28.9204
18.3128 | -35.1924 | -28.9279
18.3189 | -35.1929 | -28.9353
18.325 | -35.1934 | -28.9422
18.3311 | -35.1939 | -28.9496
18.3372 | -35.1944 | -28.9568
18.3433 | -35.195 | -28.9643
18.3494 | -35.1955 | -28.9718
18.3555 | -35.1961 | -28.9793
18.3616 | -35.1966 | -28.9866
18.3677 | -35.1972 | -28.9942
18.3738 | -35.1977 | -29.0017
18.3799 | -35.1982 | -29.009
18.386 | -35.1987 | -29.0164
18.3921 | -35.1992 | -29.0238
18.3982 | -35.1997 | -29.0308
18.4043 | -35.2002 | -29.038
18.4105 | -35.2007 | -29.0453
18.4166 | -35.2012 | -29.0526
18.4227 | -35.2017 | -29.0598
18.4288 | -35.2023 | -29.0671
18.4349 | -35.2028 | -29.0741
18.441 | -35.2034 | -29.0813
18.4471 | -35.2039 | -29.0885
18.4532 | -35.2044 | -29.0958
18.4593 | -35.2049 | -29.103
18.4654 | -35.2055 | -29.11
18.4715 | -35.206 | -29.1176
18.4776 | -35.2065 | -29.1249
18.4837 | -35.2071 | -29.1324
18.4898 | -35.2076 | -29.1398
18.4959 | -35.2081 | -29.1471
18.502 | -35.2086 | -29.1543
18.5081 | -35.2092 | -29.1616
18.5142 | -35.2097 | -29.169
18.5203 | -35.2102 | -29.1763
18.5264 | -35.2107 | -29.1836
18.5325 | -35.2113 | -29.191
18.5386 | -35.2118 | -29.198
18.5447 | -35.2123 | -29.2054
18.5508 | -35.2129 | -29.2126
18.557 | -35.2134 | -29.2201
18.5631 | -35.2139 | -29.2273
18.5692 | -35.2145 | -29.2346
18.5753 | -35.215 | -29.2417
18.5814 | -35.2156 | -29.2489
18.5875 | -35.2161 | -29.256
18.5936 | -35.2166 | -29.2632
18.5997 | -35.2171 | -29.2705
18.6058 | -35.2177 | -29.2775
18.6119 | -35.2182 | -29.2846
18.618 | -35.2187 | -29.2919
18.6241 | -35.2192 | -29.2992
18.6302 | -35.2197 | -29.3066
18.6363 | -35.2203 | -29.3138
18.6424 | -35.2208 | -29.3211
18.6485 | -35.2214 | -29.3286
18.6546 | -35.2219 | -29.3359
18.6607 | -35.2224 | -29.3433
18.6668 | -35.2229 | -29.3504
18.6729 | -35.2234 | -29.3578
18.679 | -35.224 | -29.3647
18.6851 | -35.2245 | -29.3718
18.6912 | -35.2251 | -29.379
18.6974 | -35.2256 | -29.3861
18.7035 | -35.2261 | -29.3936
18.7096 | -35.2266 | -29.4009
18.7157 | -35.2272 | -29.408
18.7218 | -35.2277 | -29.4153
18.7279 | -35.2282 | -29.4226
18.734 | -35.2287 | -29.4301
18.7401 | -35.2293 | -29.4373
18.7462 | -35.2298 | -29.4442
18.7523 | -35.2303 | -29.4514
18.7584 | -35.2309 | -29.4588
18.7645 | -35.2314 | -29.4658
18.7706 | -35.2319 | -29.4726
18.7767 | -35.2324 | -29.4799
18.7828 | -35.233 | -29.487
18.7889 | -35.2335 | -29.4943
18.795 | -35.234 | -29.5018
18.8011 | -35.2345 | -29.5091
18.8072 | -35.2351 | -29.5165
18.8133 | -35.2356 | -29.5237
18.8194 | -35.2361 | -29.5308
18.8255 | -35.2367 | -29.5381
18.8316 | -35.2372 | -29.5452
18.8377 | -35.2377 | -29.5523
18.8439 | -35.2382 | -29.5593
18.85 | -35.2387 | -29.5667
18.8561 | -35.2393 | -29.5737
18.8622 | -35.2398 | -29.5809
18.8683 | -35.2403 | -29.5879
18.8744 | -35.2409 | -29.5952
18.8805 | -35.2414 | -29.6025
18.8866 | -35.2419 | -29.6099
18.8927 | -35.2424 | -29.6168
18.8988 | -35.243 | -29.6238
18.9049 | -35.2435 | -29.631
18.911 | -35.244 | -29.6381
18.9171 | -35.2446 | -29.6452
18.9232 | -35.2451 | -29.6521
18.9293 | -35.2456 | -29.6592
18.9354 | -35.2462 | -29.6663
18.9415 | -35.2467 | -29.6734
18.9476 | -35.2472 | -29.6805
18.9537 | -35.2477 | -29.6878
18.9598 | -35.2483 | -29.6948
18.9659 | -35.2488 | -29.7021
18.972 | -35.2494 | -29.7093
18.9781 | -35.2499 | -29.7165
18.9843 | -35.2504 | -29.7238
18.9904 | -35.2509 | -29.731
18.9965 | -35.2515 | -29.7379
19.0026 | -35.252 | -29.7453
19.0087 | -35.2526 | -29.7524
19.0148 | -35.2531 | -29.7595
19.0209 | -35.2536 | -29.7668
19.027 | -35.2541 | -29.774
19.0331 | -35.2546 | -29.7807
19.0392 | -35.2551 | -29.7878
19.0453 | -35.2557 | -29.795
19.0514 | -35.2562 | -29.802
19.0575 | -35.2567 | -29.8093
19.0636 | -35.2573 | -29.816
19.0697 | -35.2578 | -29.8232
19.0758 | -35.2583 | -29.8305
19.0819 | -35.2588 | -29.8375
19.088 | -35.2594 | -29.8446
19.0941 | -35.2599 | -29.852
19.1002 | -35.2604 | -29.8589
19.1063 | -35.261 | -29.8661
19.1124 | -35.2615 | -29.8732
19.1185 | -35.262 | -29.8804
19.1246 | -35.2626 | -29.8874
19.1308 | -35.2631 | -29.8947
19.1369 | -35.2637 | -29.9016
19.143 | -35.2642 | -29.9086
19.1491 | -35.2647 | -29.9157
19.1552 | -35.2652 | -29.923
19.1613 | -35.2658 | -29.9301
19.1674 | -35.2663 | -29.9373
19.1735 | -35.2668 | -29.9442
19.1796 | -35.2673 | -29.9513
19.1857 | -35.2678 | -29.9584
19.1918 | -35.2684 | -29.9655
19.1979 | -35.2689 | -29.9726
19.204 | -35.2694 | -29.9796
19.2101 | -35.2699 | -29.9864
19.2162 | -35.2705 | -29.9934
19.2223 | -35.271 | -30.0005
19.2284 | -35.2715 | -30.0076
19.2345 | -35.272 | -30.0148
19.2406 | -35.2725 | -30.0218
19.2467 | -35.273 | -30.0289
19.2528 | -35.2736 | -30.0361
19.2589 | -35.2741 | -30.043
19.265 | -35.2746 | -30.0502
19.2712 | -35.2751 | -30.0571
19.2773 | -35.2756 | -30.064
19.2834 | -35.2762 | -30.071
19.2895 | -35.2767 | -30.0783
19.2956 | -35.2772 | -30.0853
19.3017 | -35.2777 | -30.0924
19.3078 | -35.2782 | -30.0995
19.3139 | -35.2788 | -30.1063
19.32 | -35.2793 | -30.1133
19.3261 | -35.2798 | -30.1203
19.3322 | -35.2804 | -30.1273
19.3383 | -35.2809 | -30.1342
19.3444 | -35.2814 | -30.1411
19.3505 | -35.2819 | -30.1483
19.3566 | -35.2825 | -30.1554
19.3627 | -35.283 | -30.1624
19.3688 | -35.2835 | -30.1694
19.3749 | -35.2841 | -30.1765
19.381 | -35.2846 | -30.1833
19.3871 | -35.2851 | -30.1906
19.3932 | -35.2856 | -30.1976
19.3993 | -35.2862 | -30.2045
19.4054 | -35.2867 | -30.2116
19.4115 | -35.2872 | -30.2188
19.4177 | -35.2877 | -30.2255
19.4238 | -35.2882 | -30.2325
19.4299 | -35.2888 | -30.2396
19.436 | -35.2893 | -30.2468
19.4421 | -35.2898 | -30.2539
19.4482 | -35.2904 | -30.2607
19.4543 | -35.2909 | -30.2673
19.4604 | -35.2914 | -30.2743
19.4665 | -35.292 | -30.2814
19.4726 | -35.2925 | -30.2882
19.4787 | -35.293 | -30.2951
19.4848 | -35.2935 | -30.3021
19.4909 | -35.2941 | -30.309
19.497 | -35.2946 | -30.3159
19.5031 | -35.2951 | -30.3228
19.5092 | -35.2956 | -30.3299
19.5153 | -35.2962 | -30.3367
19.5214 | -35.2967 | -30.3435
19.5275 | -35.2972 | -30.3503
19.5336 | -35.2977 | -30.3574
19.5397 | -35.2982 | -30.3644
19.5458 | -35.2987 | -30.3714
19.5519 | -35.2992 | -30.3783
19.5581 | -35.2998 | -30.385
19.5642 | -35.3003 | -30.3919
19.5703 | -35.3008 | -30.3988
19.5764 | -35.3013 | -30.4058
19.5825 | -35.3019 | -30.4127
19.5886 | -35.3024 | -30.4196
19.5947 | -35.3029 | -30.4265
19.6008 | -35.3034 | -30.4334
19.6069 | -35.304 | -30.4403
19.613 | -35.3045 | -30.4474
19.6191 | -35.305 | -30.4544
19.6252 | -35.3056 | -30.4613
19.6313 | -35.3061 | -30.4677
19.6374 | -35.3066 | -30.4748
19.6435 | -35.3072 | -30.4819
19.6496 | -35.3077 | -30.4889
19.6557 | -35.3082 | -30.4959
19.6618 | -35.3087 | -30.5021
19.6679 | -35.3093 | -30.509
19.674 | -35.3098 | -30.516
19.6801 | -35.3103 | -30.5229
19.6862 | -35.3108 | -30.5297
19.6923 | -35.3113 | -30.5368
19.6984 | -35.3119 | -30.5434
19.7046 | -35.3124 | -30.5505
19.7107 | -35.313 | -30.5572
19.7168 | -35.3135 | -30.564
19.7229 | -35.314 | -30.571
19.729 | -35.3145 | -30.5778
19.7351 | -35.315 | -30.5843
19.7412 | -35.3155 | -30.5914
19.7473 | -35.3161 | -30.5982
19.7534 | -35.3166 | -30.6048
19.7595 | -35.3171 | -30.6119
19.7656 | -35.3176 | -30.619
19.7717 | -35.3182 | -30.6256
19.7778 | -35.3187 | -30.6324
19.7839 | -35.3192 | -30.6391
19.79 | -35.3197 | -30.6458
19.7961 | -35.3203 | -30.6527
19.8022 | -35.3207 | -30.6593
19.8083 | -35.3213 | -30.6661
19.8144 | -35.3218 | -30.6728
19.8205 | -35.3224 | -30.6797
19.8266 | -35.3229 | -30.6862
19.8327 | -35.3234 | -30.6931
19.8388 | -35.3239 | -30.6998
19.845 | -35.3244 | -30.7066
19.8511 | -35.3249 | -30.7136
19.8572 | -35.3254 | -30.7205
19.8633 | -35.326 | -30.7273
19.8694 | -35.3265 | -30.7344
19.8755 | -35.327 | -30.741
19.8816 | -35.3275 | -30.7479
19.8877 | -35.328 | -30.7548
19.8938 | -35.3285 | -30.7616
19.8999 | -35.329 | -30.7685
19.906 | -35.3295 | -30.7753
19.9121 | -35.3301 | -30.7817
19.9182 | -35.3306 | -30.7885
19.9243 | -35.3311 | -30.7953
19.9304 | -35.3316 | -30.802
19.9365 | -35.3321 | -30.8088
19.9426 | -35.3326 | -30.8152
19.9487 | -35.3331 | -30.822
19.9548 | -35.3337 | -30.8287
19.9609 | -35.3342 | -30.8355
19.967 | -35.3347 | -30.8423
19.9731 | -35.3352 | -30.849
19.9792 | -35.3357 | -30.8555
19.9853 | -35.3362 | -30.8622
19.9915 | -35.3368 | -30.869
19.9976 | -35.3373 | -30.8755
20.0037 | -35.3378 | -30.8823
20.0098 | -35.3383 | -30.8891
20.0159 | -35.3388 | -30.896
20.022 | -35.3393 | -30.9028
20.0281 | -35.3398 | -30.9096
20.0342 | -35.3404 | -30.9164
20.0403 | -35.3409 | -30.9233
20.0464 | -35.3414 | -30.93
20.0525 | -35.3419 | -30.9365
20.0586 | -35.3425 | -30.9435
20.0647 | -35.343 | -30.9504
20.0708 | -35.3435 | -30.9571
20.0769 | -35.344 | -30.964
20.083 | -35.3445 | -30.9702
20.0891 | -35.345 | -30.9772
20.0952 | -35.3455 | -30.9841
20.1013 | -35.346 | -30.991
20.1074 | -35.3465 | -30.9978
20.1135 | -35.3471 | -31.0045
20.1196 | -35.3476 | -31.0107
20.1257 | -35.348 | -31.0171
20.1319 | -35.3485 | -31.0239
20.138 | -35.3491 | -31.0305
20.1441 | -35.3496 | -31.0373
20.1502 | -35.3501 | -31.0441
20.1563 | -35.3506 | -31.0507
20.1624 | -35.3512 | -31.0575
20.1685 | -35.3517 | -31.0643
20.1746 | -35.3522 | -31.071
20.1807 | -35.3527 | -31.0778
20.1868 | -35.3532 | -31.0845
20.1929 | -35.3537 | -31.0909
20.199 | -35.3542 | -31.0976
20.2051 | -35.3547 | -31.1043
20.2112 | -35.3552 | -31.1111
20.2173 | -35.3557 | -31.1177
20.2234 | -35.3562 | -31.1245
20.2295 | -35.3567 | -31.1309
20.2356 | -35.3573 | -31.1377
20.2417 | -35.3578 | -31.1445
20.2478 | -35.3583 | -31.1511
20.2539 | -35.3587 | -31.1582
20.26 | -35.3593 | -31.1646
20.2661 | -35.3598 | -31.1714
20.2723 | -35.3603 | -31.1781
20.2784 | -35.3609 | -31.1846
20.2845 | -35.3614 | -31.1913
20.2906 | -35.3619 | -31.198
20.2967 | -35.3624 | -31.2045
20.3028 | -35.3629 | -31.211
20.3089 | -35.3634 | -31.2176
20.315 | -35.3639 | -31.2244
20.3211 | -35.3645 | -31.231
20.3272 | -35.365 | -31.2379
20.3333 | -35.3655 | -31.2443
20.3394 | -35.366 | -31.251
20.3455 | -35.3665 | -31.2577
20.3516 | -35.367 | -31.2645
20.3577 | -35.3675 | -31.2711
20.3638 | -35.368 | -31.2776
20.3699 | -35.3685 | -31.2841
20.376 | -35.369 | -31.2908
20.3821 | -35.3695 | -31.2974
20.3882 | -35.37 | -31.3039
20.3943 | -35.3705 | -31.3106
20.4004 | -35.371 | -31.3169
20.4065 | -35.3715 | -31.3233
20.4126 | -35.372 | -31.33
20.4188 | -35.3725 | -31.3367
20.4249 | -35.373 | -31.3435
20.431 | -35.3736 | -31.3499
20.4371 | -35.3741 | -31.3562
20.4432 | -35.3746 | -31.3629
20.4493 | -35.3751 | -31.3696
20.4554 | -35.3756 | -31.3764
20.4615 | -35.3761 | -31.383
20.4676 | -35.3766 | -31.3897
20.4737 | -35.3771 | -31.396
20.4798 | -35.3776 | -31.4026
20.4859 | -35.3782 | -31.4095
20.492 | -35.3787 | -31.4161
20.4981 | -35.3792 | -31.4227
20.5042 | -35.3797 | -31.4293
20.5103 | -35.3802 | -31.4357
20.5164 | -35.3807 | -31.4425
20.5225 | -35.3812 | -31.4491
20.5286 | -35.3817 | -31.4558
20.5347 | -35.3822 | -31.4625
20.5408 | -35.3827 | -31.469
20.5469 | -35.3832 | -31.4758
20.553 | -35.3837 | -31.4825
20.5592 | -35.3843 | -31.4894
20.5653 | -35.3847 | -31.4959
20.5714 | -35.3852 | -31.5026
20.5775 | -35.3858 | -31.509
20.5836 | -35.3863 | -31.5156
20.5897 | -35.3868 | -31.5225
20.5958 | -35.3873 | -31.5292
20.6019 | -35.3878 | -31.5358
20.608 | -35.3883 | -31.5427
20.6141 | -35.3888 | -31.5491
20.6202 | -35.3893 | -31.5556
20.6263 | -35.3898 | -31.5622
20.6324 | -35.3903 | -31.5688
20.6385 | -35.3908 | -31.5753
20.6446 | -35.3914 | -31.582
20.6507 | -35.3919 | -31.5881
20.6568 | -35.3924 | -31.5947
20.6629 | -35.3929 | -31.6012
20.669 | -35.3934 | -31.6078
20.6751 | -35.3939 | -31.6142
20.6812 | -35.3944 | -31.6208
20.6873 | -35.3949 | -31.6277
20.6934 | -35.3954 | -31.6343
20.6995 | -35.3959 | -31.641
20.7057 | -35.3964 | -31.6475
20.7118 | -35.3969 | -31.6543
20.7179 | -35.3975 | -31.6604
20.724 | -35.398 | -31.667
20.7301 | -35.3984 | -31.6737
20.7362 | -35.3989 | -31.6803
20.7423 | -35.3994 | -31.6869
20.7484 | -35.4 | -31.6934
20.7545 | -35.4005 | -31.6994
20.7606 | -35.401 | -31.7061
20.7667 | -35.4016 | -31.7125
20.7728 | -35.402 | -31.719
20.7789 | -35.4025 | -31.7255
20.785 | -35.4031 | -31.7319
20.7911 | -35.4036 | -31.7388
20.7972 | -35.4041 | -31.7453
20.8033 | -35.4046 | -31.7522
20.8094 | -35.4051 | -31.7585
20.8155 | -35.4057 | -31.765
20.8216 | -35.4062 | -31.7716
20.8277 | -35.4067 | -31.778
20.8338 | -35.4071 | -31.7845
20.8399 | -35.4076 | -31.7912
20.8461 | -35.4082 | -31.7977
20.8522 | -35.4087 | -31.8044
20.8583 | -35.4092 | -31.8108
20.8644 | -35.4097 | -31.8177
20.8705 | -35.4102 | -31.8245
20.8766 | -35.4107 | -31.8311
20.8827 | -35.4112 | -31.8375
20.8888 | -35.4117 | -31.844
20.8949 | -35.4122 | -31.8501
20.901 | -35.4127 | -31.8567
20.9071 | -35.4132 | -31.8634
20.9132 | -35.4137 | -31.8701
20.9193 | -35.4142 | -31.8767
20.9254 | -35.4147 | -31.8835
20.9315 | -35.4152 | -31.89
20.9376 | -35.4158 | -31.8965
20.9437 | -35.4162 | -31.903
20.9498 | -35.4167 | -31.9094
20.9559 | -35.4173 | -31.9159
20.962 | -35.4178 | -31.9222
20.9681 | -35.4183 | -31.9287
20.9742 | -35.4188 | -31.9351
20.9803 | -35.4193 | -31.9415
20.9864 | -35.4199 | -31.9482
20.9926 | -35.4204 | -31.9546
20.9987 | -35.4208 | -31.9609
21.0048 | -35.4213 | -31.9674
21.0109 | -35.4218 | -31.9738
21.017 | -35.4223 | -31.9803
21.0231 | -35.4228 | -31.9866
21.0292 | -35.4233 | -31.9933
21.0353 | -35.4239 | -31.9995
21.0414 | -35.4244 | -32.006
21.0475 | -35.4249 | -32.0124
21.0536 | -35.4254 | -32.0188
21.0597 | -35.4259 | -32.0254
21.0658 | -35.4264 | -32.0319
21.0719 | -35.427 | -32.0382
21.078 | -35.4275 | -32.045
21.0841 | -35.428 | -32.0516
21.0902 | -35.4285 | -32.058
21.0963 | -35.429 | -32.0645
21.1024 | -35.4295 | -32.071
21.1085 | -35.4301 | -32.0776
21.1146 | -35.4306 | -32.0839
21.1207 | -35.4311 | -32.0903
21.1268 | -35.4316 | -32.0968
21.133 | -35.4321 | -32.1032
21.1391 | -35.4326 | -32.1093
21.1452 | -35.433 | -32.1157
21.1513 | -35.4335 | -32.1223
21.1574 | -35.4341 | -32.129
21.1635 | -35.4346 | -32.1355
21.1696 | -35.4351 | -32.1421
21.1757 | -35.4356 | -32.1482
21.1818 | -35.4361 | -32.1546
21.1879 | -35.4366 | -32.161
21.194 | -35.4372 | -32.1675
21.2001 | -35.4377 | -32.1739
21.2062 | -35.4381 | -32.1805
21.2123 | -35.4386 | -32.187
21.2184 | -35.4391 | -32.1937
21.2245 | -35.4396 | -32.2002
21.2306 | -35.4401 | -32.2065
21.2367 | -35.4406 | -32.2127
21.2428 | -35.4411 | -32.2196
21.2489 | -35.4416 | -32.2257
21.255 | -35.4421 | -32.2324
21.2611 | -35.4426 | -32.2389
21.2672 | -35.4431 | -32.2455
21.2733 | -35.4437 | -32.2519
21.2795 | -35.4443 | -32.2584
21.2856 | -35.4448 | -32.2648
21.2917 | -35.4453 | -32.2713
21.2978 | -35.4458 | -32.2779
21.3039 | -35.4463 | -32.2844
21.31 | -35.4468 | -32.291
21.3161 | -35.4474 | -32.2971
21.3222 | -35.4479 | -32.3038
21.3283 | -35.4484 | -32.3102
21.3344 | -35.4489 | -32.3167
21.3405 | -35.4494 | -32.3232
21.3466 | -35.4499 | -32.3299
21.3527 | -35.4504 | -32.3361
21.3588 | -35.4509 | -32.3427
21.3649 | -35.4514 | -32.3491
21.371 | -35.452 | -32.3556
21.3771 | -35.4524 | -32.3618
21.3832 | -35.453 | -32.3682
21.3893 | -35.4535 | -32.3747
21.3954 | -35.454 | -32.3812
21.4015 | -35.4545 | -32.3875
21.4076 | -35.455 | -32.3945
21.4137 | -35.4555 | -32.4014
21.4199 | -35.456 | -32.4076
21.426 | -35.4565 | -32.4138
21.4321 | -35.4571 | -32.4201
21.4382 | -35.4576 | -32.4265
21.4443 | -35.4581 | -32.4328
21.4504 | -35.4586 | -32.4392
21.4565 | -35.4591 | -32.4454
21.4626 | -35.4596 | -32.4519
21.4687 | -35.4601 | -32.4585
21.4748 | -35.4606 | -32.4648
21.4809 | -35.4611 | -32.471
21.487 | -35.4616 | -32.4772
21.4931 | -35.4621 | -32.4835
21.4992 | -35.4626 | -32.49
21.5053 | -35.4631 | -32.4964
21.5114 | -35.4637 | -32.503
21.5175 | -35.4642 | -32.5094
21.5236 | -35.4647 | -32.5157
21.5297 | -35.4652 | -32.5221
21.5358 | -35.4657 | -32.5284
21.5419 | -35.4662 | -32.5348
21.548 | -35.4667 | -32.5414
21.5541 | -35.4672 | -32.5481
21.5602 | -35.4677 | -32.5547
21.5664 | -35.4682 | -32.561
21.5725 | -35.4687 | -32.5674
21.5786 | -35.4692 | -32.574
21.5847 | -35.4697 | -32.5807
21.5908 | -35.4702 | -32.5872
21.5969 | -35.4707 | -32.5933
21.603 | -35.4712 | -32.6001
21.6091 | -35.4717 | -32.6065
21.6152 | -35.4722 | -32.6132
21.6213 | -35.4727 | -32.6195
21.6274 | -35.4733 | -32.6259
21.6335 | -35.4737 | -32.6322
21.6396 | -35.4743 | -32.6386
21.6457 | -35.4747 | -32.6454
21.6518 | -35.4752 | -32.6519
21.6579 | -35.4757 | -32.6583
21.664 | -35.4762 | -32.6649
21.6701 | -35.4767 | -32.6708
21.6762 | -35.4772 | -32.677
21.6823 | -35.4777 | -32.6832
21.6884 | -35.4783 | -32.6898
21.6945 | -35.4787 | -32.6962
21.7006 | -35.4793 | -32.7027
21.7068 | -35.4797 | -32.7088
21.7129 | -35.4803 | -32.7156
21.719 | -35.4808 | -32.7218
21.7251 | -35.4813 | -32.7282
21.7312 | -35.4818 | -32.7346
21.7373 | -35.4824 | -32.7408
21.7434 | -35.4829 | -32.7471
21.7495 | -35.4834 | -32.7536
21.7556 | -35.4839 | -32.76
21.7617 | -35.4844 | -32.7664
21.7678 | -35.4849 | -32.7725
21.7739 | -35.4854 | -32.7789
21.78 | -35.4859 | -32.7853
21.7861 | -35.4864 | -32.7919
21.7922 | -35.4869 | -32.7982
21.7983 | -35.4874 | -32.8046
21.8044 | -35.4879 | -32.811
21.8105 | -35.4884 | -32.8174
21.8166 | -35.4889 | -32.8241
21.8227 | -35.4895 | -32.8304
21.8288 | -35.49 | -32.8369
21.8349 | -35.4904 | -32.8433
21.841 | -35.4909 | -32.8495
21.8471 | -35.4914 | -32.8555
21.8533 | -35.4919 | -32.8621
21.8594 | -35.4925 | -32.8684
21.8655 | -35.4929 | -32.8749
21.8716 | -35.4935 | -32.8812
21.8777 | -35.494 | -32.8876
21.8838 | -35.4945 | -32.8941
21.8899 | -35.495 | -32.9006
21.896 | -35.4955 | -32.907
21.9021 | -35.496 | -32.9133
21.9082 | -35.4965 | -32.9198
21.9143 | -35.497 | -32.9261
21.9204 | -35.4975 | -32.9325
21.9265 | -35.498 | -32.9388
21.9326 | -35.4985 | -32.9453
21.9387 | -35.499 | -32.9519
21.9448 | -35.4995 | -32.9583
21.9509 | -35.5 | -32.9645
21.957 | -35.5005 | -32.9709
21.9631 | -35.501 | -32.9771
21.9692 | -35.5015 | -32.9838
21.9753 | -35.502 | -32.99
21.9814 | -35.5025 | -32.9965
21.9875 | -35.503 | -33.0027
21.9937 | -35.5035 | -33.009
21.9998 | -35.5039 | -33.0155
22.0059 | -35.5044 | -33.0218
22.012 | -35.505 | -33.0282
22.0181 | -35.5055 | -33.0344
22.0242 | -35.5059 | -33.0408
22.0303 | -35.5064 | -33.0471
22.0364 | -35.5069 | -33.0536
22.0425 | -35.5075 | -33.0597
22.0486 | -35.508 | -33.0658
22.0547 | -35.5085 | -33.0722
22.0608 | -35.509 | -33.0786
22.0669 | -35.5095 | -33.085
22.073 | -35.51 | -33.0912
22.0791 | -35.5105 | -33.0976
22.0852 | -35.511 | -33.1041
22.0913 | -35.5115 | -33.1104
22.0974 | -35.512 | -33.1166
22.1035 | -35.5125 | -33.1228
22.1096 | -35.5129 | -33.1291
22.1157 | -35.5134 | -33.1357
22.1218 | -35.5139 | -33.1421
22.1279 | -35.5144 | -33.1483
22.134 | -35.5149 | -33.1546
22.1402 | -35.5154 | -33.1612
22.1463 | -35.5159 | -33.1676
22.1524 | -35.5164 | -33.174
22.1585 | -35.5169 | -33.1802
22.1646 | -35.5174 | -33.1865
22.1707 | -35.5179 | -33.1929
22.1768 | -35.5184 | -33.1992
22.1829 | -35.5189 | -33.2055
22.189 | -35.5194 | -33.2118
22.1951 | -35.5199 | -33.2184
22.2012 | -35.5204 | -33.2247
22.2073 | -35.5209 | -33.231
22.2134 | -35.5214 | -33.2371
22.2195 | -35.5219 | -33.2433
22.2256 | -35.5224 | -33.2498
22.2317 | -35.5229 | -33.256
22.2378 | -35.5234 | -33.2626
22.2439 | -35.5239 | -33.269
22.25 | -35.5244 | -33.2753
22.2561 | -35.5249 | -33.2816
22.2622 | -35.5254 | -33.2881
22.2683 | -35.5259 | -33.2941
22.2744 | -35.5264 | -33.3005
22.2806 | -35.5269 | -33.3068
22.2867 | -35.5274 | -33.3131
22.2928 | -35.5279 | -33.3192
22.2989 | -35.5284 | -33.3255
22.305 | -35.5289 | -33.3314
22.3111 | -35.5294 | -33.3377
22.3172 | -35.5299 | -33.3441
22.3233 | -35.5304 | -33.3505
22.3294 | -35.5309 | -33.357
22.3355 | -35.5314 | -33.3636
22.3416 | -35.5319 | -33.3699
22.3477 | -35.5324 | -33.3763
22.3538 | -35.5329 | -33.3826
22.3599 | -35.5334 | -33.3888
22.366 | -35.5339 | -33.3952
22.3721 | -35.5344 | -33.4015
22.3782 | -35.5349 | -33.4077
22.3843 | -35.5354 | -33.414
22.3904 | -35.5359 | -33.4203
22.3965 | -35.5364 | -33.4267
22.4026 | -35.5369 | -33.4331
22.4087 | -35.5374 | -33.4392
22.4148 | -35.5379 | -33.4457
22.4209 | -35.5384 | -33.452
22.4271 | -35.5389 | -33.4583
22.4332 | -35.5394 | -33.4645
22.4393 | -35.5399 | -33.4707
22.4454 | -35.5404 | -33.4768
22.4515 | -35.5409 | -33.4832
22.4576 | -35.5414 | -33.4897
22.4637 | -35.5419 | -33.496
22.4698 | -35.5424 | -33.5024
22.4759 | -35.5429 | -33.5087
22.482 | -35.5434 | -33.5149
22.4881 | -35.5439 | -33.5213
22.4942 | -35.5444 | -33.528
22.5003 | -35.5449 | -33.5344
22.5064 | -35.5454 | -33.5408
22.5125 | -35.5459 | -33.547
22.5186 | -35.5464 | -33.5533
22.5247 | -35.5469 | -33.5595
22.5308 | -35.5474 | -33.5657
22.5369 | -35.5479 | -33.572
22.543 | -35.5484 | -33.5782
22.5491 | -35.5489 | -33.5842
22.5552 | -35.5494 | -33.5904
22.5613 | -35.5499 | -33.5968
22.5675 | -35.5504 | -33.6031
22.5736 | -35.5509 | -33.6093
22.5797 | -35.5514 | -33.6155
22.5858 | -35.5519 | -33.6219
22.5919 | -35.5524 | -33.628
22.598 | -35.5529 | -33.6343
22.6041 | -35.5534 | -33.6405
22.6102 | -35.5539 | -33.6469
22.6163 | -35.5544 | -33.6531
22.6224 | -35.5549 | -33.6594
22.6285 | -35.5554 | -33.6656
22.6346 | -35.5559 | -33.6719
22.6407 | -35.5564 | -33.6779
22.6468 | -35.5569 | -33.684
22.6529 | -35.5574 | -33.6901
22.659 | -35.5579 | -33.6964
22.6651 | -35.5584 | -33.7026
22.6712 | -35.559 | -33.7089
22.6773 | -35.5594 | -33.7154
22.6834 | -35.56 | -33.7218
22.6895 | -35.5605 | -33.728
22.6956 | -35.561 | -33.7342
22.7017 | -35.5615 | -33.7405
22.7079 | -35.5619 | -33.7467
22.714 | -35.5625 | -33.753
22.7201 | -35.563 | -33.7593
22.7262 | -35.5634 | -33.7657
22.7323 | -35.5639 | -33.7719
22.7384 | -35.5644 | -33.7783
22.7445 | -35.5649 | -33.7848
22.7506 | -35.5654 | -33.791
22.7567 | -35.5659 | -33.7971
22.7628 | -35.5664 | -33.8031
22.7689 | -35.5668 | -33.8093
22.775 | -35.5674 | -33.8158
22.7811 | -35.5679 | -33.8221
22.7872 | -35.5684 | -33.8284
22.7933 | -35.5689 | -33.8345
22.7994 | -35.5694 | -33.8408
22.8055 | -35.5699 | -33.8472
22.8116 | -35.5704 | -33.8532
22.8177 | -35.5709 | -33.8596
22.8238 | -35.5714 | -33.8658
22.8299 | -35.572 | -33.8719
22.836 | -35.5725 | -33.8781
22.8421 | -35.573 | -33.8846
22.8482 | -35.5734 | -33.8909
22.8544 | -35.5739 | -33.8971
22.8605 | -35.5745 | -33.9035
22.8666 | -35.575 | -33.9098
22.8727 | -35.5755 | -33.9159
22.8788 | -35.576 | -33.9219
22.8849 | -35.5764 | -33.9282
22.891 | -35.5769 | -33.9344
22.8971 | -35.5775 | -33.9406
22.9032 | -35.578 | -33.9467
22.9093 | -35.5785 | -33.953
22.9154 | -35.579 | -33.9592
22.9215 | -35.5795 | -33.9654
22.9276 | -35.5799 | -33.9717
22.9337 | -35.5804 | -33.9773
22.9398 | -35.581 | -33.9837
22.9459 | -35.5815 | -33.9901
22.952 | -35.582 | -33.9963
22.9581 | -35.5825 | -34.0027
22.9642 | -35.583 | -34.009
22.9703 | -35.5835 | -34.0149
22.9764 | -35.584 | -34.0212
22.9825 | -35.5845 | -34.0274
22.9886 | -35.585 | -34.0337
22.9948 | -35.5855 | -34.0398
23.0009 | -35.586 | -34.0462
23.007 | -35.5865 | -34.0522
23.0131 | -35.587 | -34.0584
23.0192 | -35.5875 | -34.0647
23.0253 | -35.588 | -34.0711
23.0314 | -35.5885 | -34.0775
23.0375 | -35.589 | -34.0837
23.0436 | -35.5895 | -34.0899
23.0497 | -35.59 | -34.0963
23.0558 | -35.5905 | -34.1025
23.0619 | -35.591 | -34.1089
23.068 | -35.5915 | -34.1152
23.0741 | -35.5919 | -34.1212
23.0802 | -35.5924 | -34.1275
23.0863 | -35.5929 | -34.1339
23.0924 | -35.5934 | -34.14
23.0985 | -35.5939 | -34.1462
23.1046 | -35.5944 | -34.1525
23.1107 | -35.5949 | -34.1586
23.1168 | -35.5954 | -34.1648
23.1229 | -35.5959 | -34.1709
23.129 | -35.5964 | -34.1772
23.1351 | -35.5969 | -34.1833
23.1413 | -35.5974 | -34.1897
23.1474 | -35.5979 | -34.1956
23.1535 | -35.5984 | -34.2018
23.1596 | -35.5989 | -34.2083
23.1657 | -35.5994 | -34.2145
23.1718 | -35.5999 | -34.2207
23.1779 | -35.6004 | -34.227
23.184 | -35.601 | -34.2331
23.1901 | -35.6015 | -34.2392
23.1962 | -35.602 | -34.2452
23.2023 | -35.6025 | -34.2513
23.2084 | -35.603 | -34.2575
23.2145 | -35.6035 | -34.2633
23.2206 | -35.604 | -34.2694
23.2267 | -35.6045 | -34.2758
23.2328 | -35.605 | -34.2818
23.2389 | -35.6055 | -34.2879
23.245 | -35.606 | -34.2939
23.2511 | -35.6066 | -34.2999
23.2572 | -35.6071 | -34.306
23.2633 | -35.6075 | -34.3123
23.2694 | -35.608 | -34.3185
23.2755 | -35.6085 | -34.3246
23.2817 | -35.609 | -34.3308
23.2878 | -35.6096 | -34.3367
23.2939 | -35.6101 | -34.3431
23.3 | -35.6106 | -34.3496
23.3061 | -35.611 | -34.3558
23.3122 | -35.6116 | -34.3619
23.3183 | -35.612 | -34.3679
23.3244 | -35.6125 | -34.3742
23.3305 | -35.613 | -34.3803
23.3366 | -35.6135 | -34.3865
23.3427 | -35.614 | -34.3927
23.3488 | -35.6145 | -34.3991
23.3549 | -35.615 | -34.4054
23.361 | -35.6155 | -34.4113
23.3671 | -35.616 | -34.4174
23.3732 | -35.6165 | -34.4234
23.3793 | -35.617 | -34.4294
23.3854 | -35.6175 | -34.4355
23.3915 | -35.618 | -34.4416
23.3976 | -35.6185 | -34.4476
23.4037 | -35.619 | -34.4538
23.4098 | -35.6196 | -34.4599
23.4159 | -35.62 | -34.4662
23.422 | -35.6206 | -34.4725
23.4282 | -35.6211 | -34.4784
23.4343 | -35.6216 | -34.4847
23.4404 | -35.6221 | -34.4911
23.4465 | -35.6226 | -34.4975
23.4526 | -35.6231 | -34.5039
23.4587 | -35.6236 | -34.5101
23.4648 | -35.6241 | -34.5161
23.4709 | -35.6246 | -34.5222
23.477 | -35.625 | -34.5284
23.4831 | -35.6255 | -34.5343
23.4892 | -35.626 | -34.5405
23.4953 | -35.6265 | -34.5466
23.5014 | -35.627 | -34.5527
23.5075 | -35.6275 | -34.5589
23.5136 | -35.628 | -34.5651
23.5197 | -35.6285 | -34.571
23.5258 | -35.629 | -34.5771
23.5319 | -35.6295 | -34.5829
23.538 | -35.6299 | -34.5889
23.5441 | -35.6304 | -34.595
23.5502 | -35.6309 | -34.6012
23.5563 | -35.6314 | -34.6074
23.5624 | -35.6319 | -34.6136
23.5686 | -35.6325 | -34.6199
23.5747 | -35.6329 | -34.6263
23.5808 | -35.6334 | -34.6321
23.5869 | -35.6339 | -34.6383
23.593 | -35.6344 | -34.6446
23.5991 | -35.6349 | -34.6509
23.6052 | -35.6354 | -34.6574
23.6113 | -35.6359 | -34.6637
23.6174 | -35.6364 | -34.6697
23.6235 | -35.6369 | -34.6758
23.6296 | -35.6374 | -34.6821
23.6357 | -35.6379 | -34.6883
23.6418 | -35.6384 | -34.6947
23.6479 | -35.6389 | -34.7007
23.654 | -35.6394 | -34.7066
23.6601 | -35.6399 | -34.7128
23.6662 | -35.6404 | -34.7187
23.6723 | -35.6409 | -34.7248
23.6784 | -35.6414 | -34.7309
23.6845 | -35.6419 | -34.7369
23.6906 | -35.6424 | -34.7429
23.6967 | -35.6429 | -34.749
23.7028 | -35.6434 | -34.755
23.7089 | -35.6439 | -34.761
23.7151 | -35.6444 | -34.767
23.7212 | -35.6449 | -34.773
23.7273 | -35.6454 | -34.7792
23.7334 | -35.6459 | -34.7856
23.7395 | -35.6464 | -34.7916
23.7456 | -35.6469 | -34.7976
23.7517 | -35.6474 | -34.8037
23.7578 | -35.6479 | -34.8097
23.7639 | -35.6484 | -34.8158
23.77 | -35.6489 | -34.8219
23.7761 | -35.6494 | -34.8281
23.7822 | -35.6499 | -34.8342
23.7883 | -35.6504 | -34.8399
23.7944 | -35.6508 | -34.8457
23.8005 | -35.6513 | -34.8518
23.8066 | -35.6518 | -34.858
23.8127 | -35.6523 | -34.8639
23.8188 | -35.6528 | -34.8703
23.8249 | -35.6534 | -34.8764
23.831 | -35.6539 | -34.8826
23.8371 | -35.6543 | -34.8885
23.8432 | -35.6549 | -34.8943
23.8493 | -35.6554 | -34.9004
23.8555 | -35.6558 | -34.9066
23.8616 | -35.6563 | -34.9124
23.8677 | -35.6568 | -34.9187
23.8738 | -35.6573 | -34.9245
23.8799 | -35.6579 | -34.9308
23.886 | -35.6584 | -34.9364
23.8921 | -35.6589 | -34.9422
23.8982 | -35.6594 | -34.9484
23.9043 | -35.6599 | -34.9543
23.9104 | -35.6604 | -34.9601
23.9165 | -35.6609 | -34.9662
23.9226 | -35.6614 | -34.9722
23.9287 | -35.6618 | -34.9784
23.9348 | -35.6623 | -34.9845
23.9409 | -35.6628 | -34.9905
23.947 | -35.6633 | -34.9968
23.9531 | -35.6638 | -35.0026
23.9592 | -35.6643 | -35.0089
23.9653 | -35.6647 | -35.0148
23.9714 | -35.6652 | -35.021
23.9775 | -35.6657 | -35.027
23.9836 | -35.6662 | -35.0332
23.9897 | -35.6668 | -35.039
23.9958 | -35.6672 | -35.045
24.002 | -35.6677 | -35.0512
24.0081 | -35.6682 | -35.0572
24.0142 | -35.6687 | -35.0632
24.0203 | -35.6692 | -35.0692
24.0264 | -35.6698 | -35.0753
24.0325 | -35.6703 | -35.0815
24.0386 | -35.6707 | -35.0876
24.0447 | -35.6712 | -35.0935
24.0508 | -35.6717 | -35.0995
24.0569 | -35.6722 | -35.1057
24.063 | -35.6727 | -35.1113
24.0691 | -35.6732 | -35.1172
24.0752 | -35.6738 | -35.1233
24.0813 | -35.6742 | -35.1293
24.0874 | -35.6748 | -35.1353
24.0935 | -35.6753 | -35.1414
24.0996 | -35.6757 | -35.1472
24.1057 | -35.6762 | -35.1533
24.1118 | -35.6767 | -35.1595
24.1179 | -35.6772 | -35.1658
24.124 | -35.6777 | -35.1716
24.1301 | -35.6782 | -35.1775
24.1362 | -35.6786 | -35.1836
24.1424 | -35.6792 | -35.1897
24.1485 | -35.6797 | -35.1959
24.1546 | -35.6802 | -35.202
24.1607 | -35.6807 | -35.2081
24.1668 | -35.6812 | -35.2137
24.1729 | -35.6817 | -35.2197
24.179 | -35.6822 | -35.2259
24.1851 | -35.6827 | -35.2317
24.1912 | -35.6832 | -35.2376
24.1973 | -35.6836 | -35.2434
24.2034 | -35.6841 | -35.2493
24.2095 | -35.6846 | -35.2554
24.2156 | -35.6851 | -35.2613
24.2217 | -35.6856 | -35.2673
24.2278 | -35.6861 | -35.2732
24.2339 | -35.6866 | -35.2792
24.24 | -35.6871 | -35.285
24.2461 | -35.6875 | -35.2908
24.2522 | -35.688 | -35.2969
24.2583 | -35.6885 | -35.3028
24.2644 | -35.689 | -35.3088
24.2705 | -35.6895 | -35.315
24.2766 | -35.69 | -35.3208
24.2827 | -35.6904 | -35.3267
24.2889 | -35.6909 | -35.3326
24.295 | -35.6914 | -35.3386
24.3011 | -35.6919 | -35.3446
24.3072 | -35.6923 | -35.3505
24.3133 | -35.6929 | -35.3565
24.3194 | -35.6933 | -35.3626
24.3255 | -35.6938 | -35.3685
24.3316 | -35.6943 | -35.3743
24.3377 | -35.6948 | -35.3801
24.3438 | -35.6952 | -35.3858
24.3499 | -35.6957 | -35.3919
24.356 | -35.6962 | -35.3976
24.3621 | -35.6967 | -35.4036
24.3682 | -35.6972 | -35.4097
24.3743 | -35.6976 | -35.416
24.3804 | -35.698 | -35.4219
24.3865 | -35.6985 | -35.4276
24.3926 | -35.699 | -35.4337
24.3987 | -35.6995 | -35.4396
24.4048 | -35.7 | -35.4457
24.4109 | -35.7004 | -35.4515
24.417 | -35.7009 | -35.4576
24.4231 | -35.7014 | -35.4636
24.4293 | -35.7019 | -35.4698
24.4354 | -35.7024 | -35.4757
24.4415 | -35.7029 | -35.4817
24.4476 | -35.7034 | -35.4874
24.4537 | -35.7039 | -35.4931
24.4598 | -35.7044 | -35.499
24.4659 | -35.7049 | -35.5051
24.472 | -35.7054 | -35.5109
24.4781 | -35.7059 | -35.5169
24.4842 | -35.7064 | -35.5227
24.4903 | -35.7068 | -35.5283
24.4964 | -35.7073 | -35.5341
24.5025 | -35.7078 | -35.5401
24.5086 | -35.7083 | -35.5461
24.5147 | -35.7088 | -35.5521
24.5208 | -35.7092 | -35.5577
24.5269 | -35.7097 | -35.5638
24.533 | -35.7103 | -35.5698
24.5391 | -35.7108 | -35.5757
24.5452 | -35.7113 | -35.5814
24.5513 | -35.7118 | -35.587
24.5574 | -35.7123 | -35.593
24.5635 | -35.7128 | -35.5986
24.5696 | -35.7133 | -35.6046
24.5758 | -35.7138 | -35.6106
24.5819 | -35.7143 | -35.6164
24.588 | -35.7148 | -35.6221
24.5941 | -35.7152 | -35.6279
24.6002 | -35.7157 | -35.6339
24.6063 | -35.7162 | -35.6396
24.6124 | -35.7167 | -35.6454
24.6185 | -35.7172 | -35.6513
24.6246 | -35.7177 | -35.6573
24.6307 | -35.7182 | -35.6636
24.6368 | -35.7187 | -35.6694
24.6429 | -35.7191 | -35.6753
24.649 | -35.7197 | -35.6812
24.6551 | -35.7202 | -35.6873
24.6612 | -35.7207 | -35.693
24.6673 | -35.7211 | -35.6989
24.6734 | -35.7216 | -35.7049
24.6795 | -35.7222 | -35.7106
24.6856 | -35.7226 | -35.7164
24.6917 | -35.7231 | -35.7225
24.6978 | -35.7236 | -35.7282
24.7039 | -35.7241 | -35.7342
24.71 | -35.7245 | -35.7402
24.7162 | -35.7251 | -35.7463
24.7223 | -35.7256 | -35.7524
24.7284 | -35.726 | -35.7582
24.7345 | -35.7265 | -35.7641
24.7406 | -35.727 | -35.7698
24.7467 | -35.7275 | -35.7759
24.7528 | -35.728 | -35.7816
24.7589 | -35.7285 | -35.787
24.765 | -35.7289 | -35.7926
24.7711 | -35.7294 | -35.7985
24.7772 | -35.73 | -35.8044
24.7833 | -35.7305 | -35.8101
24.7894 | -35.731 | -35.8159
24.7955 | -35.7315 | -35.8218
24.8016 | -35.732 | -35.8274
24.8077 | -35.7325 | -35.8333
24.8138 | -35.733 | -35.8395
24.8199 | -35.7334 | -35.8451
24.826 | -35.734 | -35.851
24.8321 | -35.7344 | -35.8569
24.8382 | -35.7349 | -35.8626
24.8443 | -35.7354 | -35.8684
24.8504 | -35.7359 | -35.8742
24.8565 | -35.7364 | -35.8802
24.8627 | -35.7369 | -35.8859
24.8688 | -35.7373 | -35.8917
24.8749 | -35.7378 | -35.8975
24.881 | -35.7383 | -35.9033
24.8871 | -35.7387 | -35.9091
24.8932 | -35.7392 | -35.9151
24.8993 | -35.7397 | -35.9208
24.9054 | -35.7402 | -35.9263
24.9115 | -35.7407 | -35.9322
24.9176 | -35.7412 | -35.938
24.9237 | -35.7417 | -35.944
24.9298 | -35.7422 | -35.95
24.9359 | -35.7427 | -35.9558
24.942 | -35.7432 | -35.9612
24.9481 | -35.7437 | -35.967
24.9542 | -35.7441 | -35.9726
24.9603 | -35.7447 | -35.9785
24.9664 | -35.7452 | -35.9842
24.9725 | -35.7457 | -35.9902
24.9786 | -35.7462 | -35.9957
24.9847 | -35.7466 | -36.0014
24.9908 | -35.7471 | -36.0072
24.9969 | -35.7476 | -36.0131
25.0031 | -35.7481 | -36.0191
25.0092 | -35.7485 | -36.0249
25.0153 | -35.749 | -36.0305
25.0214 | -35.7495 | -36.0363
25.0275 | -35.7499 | -36.042
25.0336 | -35.7504 | -36.0477
25.0397 | -35.7509 | -36.0536
25.0458 | -35.7514 | -36.0593
25.0519 | -35.7519 | -36.0651
25.058 | -35.7523 | -36.071
25.0641 | -35.7528 | -36.0766
25.0702 | -35.7533 | -36.0823
25.0763 | -35.7538 | -36.0882
25.0824 | -35.7543 | -36.094
25.0885 | -35.7547 | -36.0996
25.0946 | -35.7552 | -36.1054
25.1007 | -35.7557 | -36.1112
25.1068 | -35.7562 | -36.1167
25.1129 | -35.7566 | -36.1227
25.119 | -35.7571 | -36.1285
25.1251 | -35.7576 | -36.1341
25.1312 | -35.7581 | -36.1405
25.1373 | -35.7586 | -36.1465
25.1435 | -35.7591 | -36.1523
25.1496 | -35.7596 | -36.1579
25.1557 | -35.7601 | -36.1637
25.1618 | -35.7606 | -36.1695
25.1679 | -35.7611 | -36.1754
25.174 | -35.7616 | -36.1812
25.1801 | -35.7621 | -36.1871
25.1862 | -35.7625 | -36.1927
25.1923 | -35.763 | -36.1986
25.1984 | -35.7635 | -36.2044
25.2045 | -35.7639 | -36.2102
25.2106 | -35.7644 | -36.2159
25.2167 | -35.7649 | -36.2218
25.2228 | -35.7653 | -36.2276
25.2289 | -35.7658 | -36.2334
25.235 | -35.7663 | -36.2395
25.2411 | -35.7668 | -36.245
25.2472 | -35.7673 | -36.2508
25.2533 | -35.7678 | -36.2565
25.2594 | -35.7682 | -36.262
25.2655 | -35.7687 | -36.2677
25.2716 | -35.7692 | -36.2735
25.2777 | -35.7697 | -36.2794
25.2838 | -35.7702 | -36.2853
25.29 | -35.7707 | -36.2908
25.2961 | -35.7711 | -36.2962
25.3022 | -35.7716 | -36.302
25.3083 | -35.7721 | -36.3078
25.3144 | -35.7726 | -36.3137
25.3205 | -35.773 | -36.3195
25.3266 | -35.7735 | -36.3252
25.3327 | -35.774 | -36.3311
25.3388 | -35.7745 | -36.337
25.3449 | -35.7749 | -36.3428
25.351 | -35.7754 | -36.3486
25.3571 | -35.7759 | -36.3544
25.3632 | -35.7763 | -36.3602
25.3693 | -35.7769 | -36.366
25.3754 | -35.7773 | -36.3715
25.3815 | -35.7778 | -36.3773
25.3876 | -35.7783 | -36.383
25.3937 | -35.7788 | -36.3887
25.3998 | -35.7792 | -36.3942
25.4059 | -35.7797 | -36.4001
25.412 | -35.7802 | -36.406
25.4181 | -35.7806 | -36.4116
25.4242 | -35.7811 | -36.4171
25.4304 | -35.7816 | -36.4229
25.4365 | -35.7821 | -36.4285
25.4426 | -35.7825 | -36.4342
25.4487 | -35.783 | -36.4399
25.4548 | -35.7835 | -36.4457
25.4609 | -35.7839 | -36.4514
25.467 | -35.7844 | -36.4569
25.4731 | -35.7849 | -36.4627
25.4792 | -35.7854 | -36.4683
25.4853 | -35.7859 | -36.4742
25.4914 | -35.7864 | -36.4798
25.4975 | -35.7869 | -36.4854
25.5036 | -35.7873 | -36.4909
25.5097 | -35.7878 | -36.4965
25.5158 | -35.7883 | -36.5022
25.5219 | -35.7887 | -36.5078
25.528 | -35.7892 | -36.5136
25.5341 | -35.7897 | -36.5193
25.5402 | -35.7901 | -36.5247
25.5463 | -35.7907 | -36.5303
25.5524 | -35.7911 | -36.536
25.5585 | -35.7916 | -36.5417
25.5646 | -35.7921 | -36.5475
25.5707 | -35.7926 | -36.5531
25.5769 | -35.793 | -36.5588
25.583 | -35.7935 | -36.5645
25.5891 | -35.794 | -36.5701
25.5952 | -35.7944 | -36.5758
25.6013 | -35.7949 | -36.5814
25.6074 | -35.7954 | -36.5867
25.6135 | -35.7958 | -36.5925
25.6196 | -35.7963 | -36.5984
25.6257 | -35.7968 | -36.6041
25.6318 | -35.7973 | -36.6097
25.6379 | -35.7977 | -36.6154
25.644 | -35.7982 | -36.621
25.6501 | -35.7987 | -36.6268
25.6562 | -35.7991 | -36.6324
25.6623 | -35.7996 | -36.6383
25.6684 | -35.8 | -36.6442
25.6745 | -35.8006 | -36.6499
25.6806 | -35.801 | -36.6553
25.6867 | -35.8015 | -36.661
25.6928 | -35.802 | -36.6667
25.6989 | -35.8025 | -36.6726
25.705 | -35.8029 | -36.678
25.7111 | -35.8034 | -36.6841
25.7173 | -35.8039 | -36.6897
25.7234 | -35.8044 | -36.6952
25.7295 | -35.8048 | -36.7008
25.7356 | -35.8053 | -36.7067
25.7417 | -35.8058 | -36.7123
25.7478 | -35.8062 | -36.7174
25.7539 | -35.8066 | -36.7231
25.76 | -35.8071 | -36.7285
25.7661 | -35.8076 | -36.7337
25.7722 | -35.8081 | -36.7394
25.7783 | -35.8085 | -36.7452
25.7844 | -35.8089 | -36.7505
25.7905 | -35.8094 | -36.7562
25.7966 | -35.8098 | -36.7618
25.8027 | -35.8102 | -36.7674
25.8088 | -35.8107 | -36.7727
25.8149 | -35.8111 | -36.7779
25.821 | -35.8116 | -36.7833
25.8271 | -35.812 | -36.7888
25.8332 | -35.8125 | -36.7944
25.8393 | -35.8129 | -36.8002
25.8454 | -35.8134 | -36.8059
25.8515 | -35.8139 | -36.8115
25.8576 | -35.8144 | -36.8169
25.8638 | -35.8148 | -36.8227
25.8699 | -35.8152 | -36.8285
25.876 | -35.8157 | -36.8344
25.8821 | -35.8162 | -36.84
25.8882 | -35.8166 | -36.8453
25.8943 | -35.8171 | -36.8513
25.9004 | -35.8175 | -36.8568
25.9065 | -35.818 | -36.8625
25.9126 | -35.8186 | -36.8679
25.9187 | -35.819 | -36.8733
25.9248 | -35.8195 | -36.8788
25.9309 | -35.8199 | -36.8844
25.937 | -35.8204 | -36.8901
25.9431 | -35.8209 | -36.8958
25.9492 | -35.8213 | -36.9011
25.9553 | -35.8218 | -36.9069
25.9614 | -35.8222 | -36.912
25.9675 | -35.8226 | -36.9179
25.9736 | -35.8231 | -36.9238
25.9797 | -35.8235 | -36.9294
25.9858 | -35.824 | -36.9352
25.9919 | -35.8244 | -36.9413
25.998 | -35.8249 | -36.9467
26.0042 | -35.8253 | -36.9525
26.0103 | -35.8257 | -36.9582
26.0164 | -35.8262 | -36.9636
26.0225 | -35.8267 | -36.9688
26.0286 | -35.8271 | -36.9743
26.0347 | -35.8276 | -36.9793
26.0408 | -35.828 | -36.9849
26.0469 | -35.8284 | -36.9903
26.053 | -35.8289 | -36.9959
26.0591 | -35.8293 | -37.0015
26.0652 | -35.8298 | -37.0072
26.0713 | -35.8303 | -37.0125
26.0774 | -35.8307 | -37.0181
26.0835 | -35.8312 | -37.0239
26.0896 | -35.8316 | -37.0294
26.0957 | -35.8321 | -37.0349
26.1018 | -35.8324 | -37.0399
26.1079 | -35.8329 | -37.0456
26.114 | -35.8333 | -37.0515
26.1201 | -35.8338 | -37.057
26.1262 | -35.8343 | -37.0627
26.1323 | -35.8348 | -37.0681
26.1384 | -35.8353 | -37.0732
26.1445 | -35.8357 | -37.0789
26.1507 | -35.8362 | -37.0847
26.1568 | -35.8367 | -37.0901
26.1629 | -35.8372 | -37.0956
26.169 | -35.8377 | -37.1012
26.1751 | -35.8381 | -37.1068
26.1812 | -35.8385 | -37.1124
26.1873 | -35.8389 | -37.1179
26.1934 | -35.8394 | -37.1235
26.1995 | -35.8399 | -37.1291
26.2056 | -35.8403 | -37.1346
26.2117 | -35.8407 | -37.1401
26.2178 | -35.8412 | -37.1458
26.2239 | -35.8416 | -37.1513
26.23 | -35.8421 | -37.1567
26.2361 | -35.8425 | -37.1625
26.2422 | -35.843 | -37.1679
26.2483 | -35.8435 | -37.1736
26.2544 | -35.8439 | -37.1788
26.2605 | -35.8444 | -37.1843
26.2666 | -35.8448 | -37.1898
26.2727 | -35.8453 | -37.1954
26.2788 | -35.8457 | -37.2008
26.2849 | -35.8462 | -37.2064
26.2911 | -35.8467 | -37.2118
26.2972 | -35.8471 | -37.2175
26.3033 | -35.8476 | -37.223
26.3094 | -35.848 | -37.2284
26.3155 | -35.8485 | -37.2333
26.3216 | -35.8489 | -37.2387
26.3277 | -35.8494 | -37.2438
26.3338 | -35.8498 | -37.2492
26.3399 | -35.8503 | -37.255
26.346 | -35.8508 | -37.2605
26.3521 | -35.8512 | -37.266
26.3582 | -35.8516 | -37.2713
26.3643 | -35.8521 | -37.2767
26.3704 | -35.8526 | -37.2821
26.3765 | -35.853 | -37.2874
26.3826 | -35.8534 | -37.2926
26.3887 | -35.8539 | -37.2982
26.3948 | -35.8543 | -37.3039
26.4009 | -35.8548 | -37.3095
26.407 | -35.8552 | -37.3149
26.4131 | -35.8557 | -37.3206
26.4192 | -35.8561 | -37.3257
26.4253 | -35.8566 | -37.3313
26.4314 | -35.857 | -37.3371
26.4376 | -35.8575 | -37.3424
26.4437 | -35.858 | -37.3479
26.4498 | -35.8584 | -37.3535
26.4559 | -35.8588 | -37.3587
26.462 | -35.8592 | -37.364
26.4681 | -35.8597 | -37.3696
26.4742 | -35.8602 | -37.3754
26.4803 | -35.8606 | -37.3807
26.4864 | -35.8611 | -37.3859
26.4925 | -35.8616 | -37.3912
26.4986 | -35.862 | -37.3963
26.5047 | -35.8625 | -37.4017
26.5108 | -35.8629 | -37.4069
26.5169 | -35.8634 | -37.4125
26.523 | -35.8639 | -37.4176
26.5291 | -35.8643 | -37.423
26.5352 | -35.8647 | -37.4285
26.5413 | -35.8652 | -37.4338
26.5474 | -35.8656 | -37.4393
26.5535 | -35.866 | -37.4446
26.5596 | -35.8665 | -37.4503
26.5657 | -35.867 | -37.4559
26.5718 | -35.8674 | -37.4611
26.578 | -35.8678 | -37.4667
26.5841 | -35.8683 | -37.472
26.5902 | -35.8687 | -37.4773
26.5963 | -35.8692 | -37.4825
26.6024 | -35.8697 | -37.488
26.6085 | -35.8701 | -37.4936
26.6146 | -35.8706 | -37.4989
26.6207 | -35.871 | -37.5044
26.6268 | -35.8714 | -37.5099
26.6329 | -35.8718 | -37.515
26.639 | -35.8722 | -37.5201
26.6451 | -35.8726 | -37.5258
26.6512 | -35.873 | -37.5307
26.6573 | -35.8735 | -37.5364
26.6634 | -35.8739 | -37.5418
26.6695 | -35.8743 | -37.5478
26.6756 | -35.8748 | -37.5532
26.6817 | -35.8752 | -37.5587
26.6878 | -35.8756 | -37.5641
26.6939 | -35.876 | -37.5694
26.7 | -35.8765 | -37.5747
26.7061 | -35.8769 | -37.5799
26.7122 | -35.8773 | -37.5854
26.7183 | -35.8778 | -37.5906
26.7245 | -35.8782 | -37.5959
26.7306 | -35.8787 | -37.6013
26.7367 | -35.8791 | -37.6063
26.7428 | -35.8795 | -37.6119
26.7489 | -35.88 | -37.6171
26.755 | -35.8805 | -37.6226
26.7611 | -35.8808 | -37.6286
26.7672 | -35.8813 | -37.6348
26.7733 | -35.8816 | -37.64
26.7794 | -35.8821 | -37.6456
26.7855 | -35.8825 | -37.651
26.7916 | -35.883 | -37.6563
26.7977 | -35.8835 | -37.6619
26.8038 | -35.8839 | -37.667
26.8099 | -35.8843 | -37.6725
26.816 | -35.8847 | -37.6779
26.8221 | -35.8852 | -37.6834
26.8282 | -35.8856 | -37.6887
26.8343 | -35.886 | -37.6941
26.8404 | -35.8864 | -37.699
26.8465 | -35.8868 | -37.7046
26.8526 | -35.8873 | -37.7098
26.8587 | -35.8877 | -37.7155
26.8649 | -35.8882 | -37.7212
26.871 | -35.8886 | -37.7264
26.8771 | -35.889 | -37.7319
26.8832 | -35.8895 | -37.7373
26.8893 | -35.8899 | -37.7425
26.8954 | -35.8903 | -37.7479
26.9015 | -35.8907 | -37.7534
26.9076 | -35.8912 | -37.759
26.9137 | -35.8916 | -37.764
26.9198 | -35.8921 | -37.7693
26.9259 | -35.8925 | -37.7744
26.932 | -35.8929 | -37.7795
26.9381 | -35.8934 | -37.7848
26.9442 | -35.8938 | -37.7901
26.9503 | -35.8942 | -37.7953
26.9564 | -35.8946 | -37.8002
26.9625 | -35.895 | -37.8059
26.9686 | -35.8954 | -37.8111
26.9747 | -35.8959 | -37.8162
26.9808 | -35.8963 | -37.8214
26.9869 | -35.8967 | -37.8268
26.993 | -35.8971 | -37.8323
26.9991 | -35.8976 | -37.8374
27.0052 | -35.898 | -37.8426
27.0114 | -35.8984 | -37.8477
27.0175 | -35.8989 | -37.8529
27.0236 | -35.8993 | -37.8584
27.0297 | -35.8998 | -37.8637
27.0358 | -35.9002 | -37.8692
27.0419 | -35.9007 | -37.8745
27.048 | -35.9012 | -37.8798
27.0541 | -35.9016 | -37.8849
27.0602 | -35.902 | -37.8899
27.0663 | -35.9023 | -37.8954
27.0724 | -35.9028 | -37.9009
27.0785 | -35.9032 | -37.9069
27.0846 | -35.9036 | -37.912
27.0907 | -35.9041 | -37.9173
27.0968 | -35.9045 | -37.9227
27.1029 | -35.905 | -37.9278
27.109 | -35.9054 | -37.9331
27.1151 | -35.9058 | -37.9387
27.1212 | -35.9061 | -37.9439
27.1273 | -35.9066 | -37.9491
27.1334 | -35.907 | -37.9546
27.1395 | -35.9074 | -37.9599
27.1456 | -35.9078 | -37.9651
27.1518 | -35.9082 | -37.9705
27.1579 | -35.9086 | -37.9752
27.164 | -35.909 | -37.9806
27.1701 | -35.9094 | -37.9857
27.1762 | -35.9099 | -37.9911
27.1823 | -35.9103 | -37.9964
27.1884 | -35.9107 | -38.0019
27.1945 | -35.9111 | -38.007
27.2006 | -35.9116 | -38.0123
27.2067 | -35.912 | -38.0178
27.2128 | -35.9124 | -38.0228
27.2189 | -35.9127 | -38.0279
27.225 | -35.9132 | -38.0336
27.2311 | -35.9135 | -38.0386
27.2372 | -35.914 | -38.044
27.2433 | -35.9144 | -38.049
27.2494 | -35.9149 | -38.0543
27.2555 | -35.9154 | -38.0598
27.2616 | -35.9157 | -38.0647
27.2677 | -35.9161 | -38.07
27.2738 | -35.9165 | -38.075
27.2799 | -35.917 | -38.0801
27.286 | -35.9174 | -38.0852
27.2921 | -35.9178 | -38.0906
27.2983 | -35.9181 | -38.0957
27.3044 | -35.9186 | -38.1012
27.3105 | -35.919 | -38.1064
27.3166 | -35.9194 | -38.1116
27.3227 | -35.9198 | -38.1167
27.3288 | -35.9202 | -38.122
27.3349 | -35.9206 | -38.127
27.341 | -35.921 | -38.1323
27.3471 | -35.9215 | -38.1373
27.3532 | -35.9219 | -38.1425
27.3593 | -35.9223 | -38.148
27.3654 | -35.9228 | -38.1536
27.3715 | -35.9231 | -38.1582
27.3776 | -35.9234 | -38.1638
27.3837 | -35.9238 | -38.169
27.3898 | -35.9243 | -38.1744
27.3959 | -35.9247 | -38.18
27.402 | -35.9251 | -38.1848
27.4081 | -35.9255 | -38.1899
27.4142 | -35.9259 | -38.1952
27.4203 | -35.9263 | -38.2004
27.4264 | -35.9267 | -38.2053
27.4325 | -35.9272 | -38.2108
27.4387 | -35.9275 | -38.2157
27.4448 | -35.9279 | -38.221
27.4509 | -35.9282 | -38.226
27.457 | -35.9287 | -38.2313
27.4631 | -35.9291 | -38.2364
27.4692 | -35.9295 | -38.2417
27.4753 | -35.9299 | -38.2472
27.4814 | -35.9303 | -38.2519
27.4875 | -35.9307 | -38.2568
27.4936 | -35.9311 | -38.262
27.4997 | -35.9315 | -38.2671
27.5058 | -35.932 | -38.2722
27.5119 | -35.9323 | -38.2772
27.518 | -35.9327 | -38.2822
27.5241 | -35.9331 | -38.2876
27.5302 | -35.9335 | -38.2927
27.5363 | -35.934 | -38.2975
27.5424 | -35.9343 | -38.3027
27.5485 | -35.9347 | -38.3077
27.5546 | -35.9351 | -38.3129
27.5607 | -35.9355 | -38.3182
27.5668 | -35.9359 | -38.3232
27.5729 | -35.9364 | -38.3283
27.5791 | -35.9367 | -38.3334
27.5852 | -35.9372 | -38.3388
27.5913 | -35.9375 | -38.3444
27.5974 | -35.938 | -38.3499
27.6035 | -35.9384 | -38.3548
27.6096 | -35.9388 | -38.3599
27.6157 | -35.9392 | -38.3652
27.6218 | -35.9396 | -38.3703
27.6279 | -35.94 | -38.3755
27.634 | -35.9404 | -38.3808
27.6401 | -35.9408 | -38.3859
27.6462 | -35.9412 | -38.3912
27.6523 | -35.9416 | -38.3959
27.6584 | -35.942 | -38.401
27.6645 | -35.9424 | -38.4067
27.6706 | -35.9428 | -38.4123
27.6767 | -35.9431 | -38.4173
27.6828 | -35.9435 | -38.4226
27.6889 | -35.9439 | -38.428
27.695 | -35.9443 | -38.4336
27.7011 | -35.9446 | -38.4386
27.7072 | -35.945 | -38.4437
27.7133 | -35.9454 | -38.4488
27.7194 | -35.9458 | -38.4536
27.7256 | -35.9462 | -38.4586
27.7317 | -35.9466 | -38.4638
27.7378 | -35.9469 | -38.4693
27.7439 | -35.9473 | -38.4747
27.75 | -35.9478 | -38.4795
27.7561 | -35.9482 | -38.4845
27.7622 | -35.9486 | -38.4897
27.7683 | -35.949 | -38.4948
27.7744 | -35.9494 | -38.4998
27.7805 | -35.9497 | -38.505
27.7866 | -35.9502 | -38.5102
27.7927 | -35.9506 | -38.5149
27.7988 | -35.951 | -38.5203
27.8049 | -35.9513 | -38.5254
27.811 | -35.9517 | -38.5305
27.8171 | -35.9521 | -38.5358
27.8232 | -35.9525 | -38.5408
27.8293 | -35.9529 | -38.5461
27.8354 | -35.9533 | -38.5512
27.8415 | -35.9537 | -38.5563
27.8476 | -35.9541 | -38.5618
27.8537 | -35.9545 | -38.5671
27.8598 | -35.9548 | -38.5723
27.866 | -35.9551 | -38.5776
27.8721 | -35.9555 | -38.5826
27.8782 | -35.9559 | -38.5879
27.8843 | -35.9563 | -38.5931
27.8904 | -35.9567 | -38.599
27.8965 | -35.957 | -38.6036
27.9026 | -35.9574 | -38.6085
27.9087 | -35.9577 | -38.6136
27.9148 | -35.9581 | -38.6188
27.9209 | -35.9585 | -38.6237
27.927 | -35.959 | -38.629
27.9331 | -35.9593 | -38.6342
27.9392 | -35.9597 | -38.6394
27.9453 | -35.9601 | -38.6444
27.9514 | -35.9604 | -38.6493
27.9575 | -35.9608 | -38.6548
27.9636 | -35.9612 | -38.6603
27.9697 | -35.9617 | -38.6647
27.9758 | -35.962 | -38.6698
27.9819 | -35.9624 | -38.675
27.988 | -35.9628 | -38.6804
27.9941 | -35.9631 | -38.6853
28.0002 | -35.9634 | -38.6903
28.0063 | -35.9638 | -38.6957
28.0125 | -35.9641 | -38.7007
28.0186 | -35.9646 | -38.7057
28.0247 | -35.965 | -38.7109
28.0308 | -35.9654 | -38.7157
28.0369 | -35.9659 | -38.7211
28.043 | -35.9662 | -38.7261
28.0491 | -35.9666 | -38.7314
28.0552 | -35.967 | -38.7365
28.0613 | -35.9674 | -38.742
28.0674 | -35.9678 | -38.7467
28.0735 | -35.9682 | -38.7513

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.6486 | -42.9438
21.4908 | -35.6482 | -35.9707
16.4948 | -34.9812 | -28.4138
12.6602 | -34.5894 | -21.9949
9.7171 | -34.3569 | -16.9754
7.4582 | -34.1988 | -12.9068
5.7244 | -34.0857 | -9.6097
4.3936 | -34.0164 | -6.952
3.3722 | -33.9922 | -4.8574
2.5883 | -33.9835 | -3.2688
1.9866 | -33.9792 | -1.9971
1.5248 | -33.9799 | -0.98454
1.1703 | -33.9971 | -0.31673
1 | -34.0076 | 0.010305
0.89822 | -33.995 | 0.28702
0.68941 | -33.974 | 0.87544
0.52919 | -33.9736 | 1.3313
0.40612 | -33.9772 | 1.9444
0.31171 | -33.9728 | 2.3989
0.23928 | -33.9977 | 3.1198
0.18358 | -33.9951 | 3.5294
0.14091 | -33.9883 | 4.4148
0.10819 | -34.0126 | 5.5746
0.083017 | -34.051 | 7.0421
0.063705 | -34.0738 | 8.9173
0.048923 | -34.143 | 11.4098
0.037479 | -34.248 | 14.6006
0.028801 | -34.42 | 18.6703
0.022078 | -34.7367 | 24.0928
0.017023 | -35.2272 | 30.4959
0.013018 | -36.0333 | 38.1642
0.010014 | -37.144 | 46.098

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003666 | -28.3558
0.007331 | -34.4618
0.010997 | -37.7909
0.014663 | -40.7841
0.018328 | -42.5064
0.021994 | -44.2821
0.02566 | -45.6622
0.029326 | -48.0161
0.032991 | -49.7561
0.036657 | -50.8232
0.040323 | -52.2897
0.043988 | -53.4397
0.047654 | -54.0718
0.05132 | -55.4921
0.054985 | -55.9656
0.058651 | -56.573
0.062317 | -57.115
0.065982 | -57.5302
0.069648 | -58.3692
0.073314 | -58.9762
0.076979 | -59.6579
0.080645 | -60.9373
0.084311 | -60.9624
0.087977 | -61.1445
0.091642 | -62.1584
0.095308 | -61.7609
0.098974 | -62.2277
0.10264 | -62.5358
0.1063 | -62.5936
0.10997 | -63.4822
0.11364 | -64.9927
0.1173 | -65.1335
0.12097 | -65.4134
0.12463 | -65.8528
0.1283 | -65.9043
0.13196 | -65.5214
0.13563 | -66.3665
0.1393 | -67.3616
0.14296 | -67.553
0.14663 | -66.216
0.15029 | -66.2897
0.15396 | -67.1799
0.15762 | -67.901
0.16129 | -68.5078
0.16496 | -67.8562
0.16862 | -67.2433
0.17229 | -67.832
0.17595 | -68.0348
0.17962 | -69.1521
0.18328 | -69.5231
0.18695 | -70.3548
0.19062 | -70.57
0.19428 | -69.9138
0.19795 | -70.0198
0.20161 | -69.5607
0.20528 | -69.9092
0.20894 | -70.7769
0.21261 | -72.7203
0.21628 | -71.7401
0.21994 | -71.759
0.22361 | -72.1782
0.22727 | -72.4904
0.23094 | -71.9349
0.2346 | -72.0512
0.23827 | -71.7863
0.24194 | -72.4566
0.2456 | -73.174
0.24927 | -73.2946
0.25293 | -72.7237
0.2566 | -73.3202
0.26026 | -73.5604
0.26393 | -72.7388
0.2676 | -73.3753
0.27126 | -74.4284
0.27493 | -74.6268
0.27859 | -73.4062
0.28226 | -72.5566
0.28592 | -73.7898
0.28959 | -75.0427
0.29325 | -74.7736
0.29692 | -74.2689
0.30059 | -74.7046
0.30425 | -74.2395
0.30792 | -74.8221
0.31158 | -74.9939
0.31525 | -75.8387
0.31892 | -76.284
0.32258 | -76.2846
0.32625 | -76.2146
0.32991 | -75.6854
0.33358 | -76.4625
0.33724 | -75.7073
0.34091 | -76.4602
0.34458 | -77.2039
0.34824 | -75.8768
0.35191 | -76.4114
0.35557 | -76.2226
0.35924 | -76.7807
0.3629 | -76.7238
0.36657 | -76.0972
0.37023 | -77.0361
0.3739 | -76.9791
0.37757 | -77.5787
0.38123 | -78.8343
0.3849 | -78.6117
0.38856 | -78.451
0.39223 | -78.121
0.39589 | -77.4237
0.39956 | -77.1458
0.40323 | -77.269
0.40689 | -77.1999
0.41056 | -77.8611
0.41422 | -77.5439
0.41789 | -78.1463
0.42155 | -78.6275
0.42522 | -79.4046
0.42889 | -78.8655
0.43255 | -78.0316
0.43622 | -77.9197
0.43988 | -78.1962
0.44355 | -77.936
0.44721 | -77.69
0.45088 | -78.716
0.45454 | -79.8376
0.45821 | -79.1034
0.46188 | -78.5589
0.46554 | -79.3701
0.46921 | -80.0335
0.47287 | -79.2013
0.47654 | -78.9677
0.4802 | -79.7858
0.48387 | -79.9989
0.48754 | -79.9638
0.4912 | -79.8963
0.49487 | -80.0316
0.49853 | -80.4444
0.5022 | -80.4131
0.50587 | -80.3016
0.50953 | -79.1262
0.5132 | -78.5225
0.51686 | -79.4103
0.52053 | -79.1986
0.52419 | -79.8495
0.52786 | -79.593
0.53153 | -79.4837
0.53519 | -79.9692
0.53886 | -80.6575
0.54252 | -80.7649
0.54619 | -80.3915
0.54985 | -80.3706
0.55352 | -80.3848
0.55719 | -80.1609
0.56085 | -80.16
0.56452 | -80.3847
0.56818 | -80.5535
0.57185 | -80.7465
0.57551 | -81.146
0.57918 | -81.6829
0.58284 | -81.4527
0.58651 | -81.2224
0.59018 | -81.6779
0.59384 | -82.4027
0.59751 | -82.2414
0.60117 | -81.4119
0.60484 | -81.1734
0.6085 | -81.5245
0.61217 | -81.8149
0.61584 | -81.6686
0.6195 | -81.5953
0.62317 | -81.8402
0.62683 | -81.902
0.6305 | -81.8128
0.63416 | -81.7086
0.63783 | -81.6987
0.6415 | -81.8923
0.64516 | -82.3693
0.64883 | -82.6763
0.65249 | -82.8338
0.65616 | -82.8903
0.65982 | -82.8333
0.66349 | -83.1466
0.66716 | -83.611
0.67082 | -83.6175
0.67449 | -83.1164
0.67815 | -82.8838
0.68182 | -82.9371
0.68548 | -83.3125
0.68915 | -83.7819
0.69281 | -84.0434
0.69648 | -84.0531
0.70015 | -83.2953
0.70381 | -82.4174
0.70748 | -82.3877
0.71114 | -83.0143
0.71481 | -83.3889
0.71847 | -83.4135
0.72214 | -83.2724
0.72581 | -83.087
0.72947 | -83.171
0.73314 | -83.4416
0.7368 | -83.5258
0.74047 | -83.3038
0.74413 | -83.3253
0.7478 | -83.6761
0.75147 | -84.1086
0.75513 | -84.0153
0.7588 | -83.3365
0.76246 | -83.2408
0.76613 | -83.3841
0.7698 | -83.1615
0.77346 | -83.1307
0.77713 | -83.1994
0.78079 | -83.3798
0.78446 | -83.9165
0.78812 | -84.1368
0.79179 | -84.0572
0.79546 | -83.8058
0.79912 | -83.4819
0.80279 | -83.1933
0.80645 | -83.2165
0.81012 | -84.3651
0.81378 | -85.3264
0.81745 | -85.0801
0.82111 | -85.1641
0.82478 | -85.5853
0.82845 | -85.5471
0.83211 | -84.8889
0.83578 | -84.3392
0.83944 | -84.3467
0.84311 | -84.6911
0.84677 | -84.696
0.85044 | -84.6712
0.85411 | -84.7212
0.85777 | -84.589
0.86144 | -84.4955
0.8651 | -84.4617
0.86877 | -84.5904
0.87243 | -84.864
0.8761 | -85.0473
0.87977 | -85.285
0.88343 | -85.7149
0.8871 | -85.6368
0.89076 | -85.1624
0.89443 | -85.0452
0.89809 | -85.5161
0.90176 | -85.9579
0.90543 | -86.0665
0.90909 | -86.0939
0.91276 | -85.9202
0.91642 | -86.022
0.92009 | -86.7794
0.92375 | -87.1812
0.92742 | -86.622
0.93109 | -85.9435
0.93475 | -85.7261
0.93842 | -85.6106
0.94208 | -85.3314
0.94575 | -85.3706
0.94941 | -85.5748
0.95308 | -85.4982
0.95674 | -85.4401
0.96041 | -85.6927
0.96408 | -86.1492
0.96774 | -86.474
0.97141 | -86.5198
0.97507 | -86.3184
0.97874 | -86.4535
0.9824 | -86.7809
0.98607 | -86.8085
0.98974 | -86.3373
0.9934 | -86.058
0.99707 | -86.1196
1.0007 | -86.1768
1.0044 | -86.4358
1.0081 | -86.8646
1.0117 | -86.9215
1.0154 | -87.1089
1.0191 | -87.0593
1.0227 | -86.879
1.0264 | -87.0134
1.0301 | -87.1336
1.0337 | -87.3497
1.0374 | -88.0121
1.0411 | -88.0727
1.0447 | -87.0971
1.0484 | -86.2734
1.0521 | -86.3005
1.0557 | -86.7748
1.0594 | -87.0425
1.0631 | -87.1554
1.0667 | -87.2321
1.0704 | -87.3106
1.074 | -87.3209
1.0777 | -87.3073
1.0814 | -87.3296
1.085 | -87.3227
1.0887 | -87.2468
1.0924 | -87.1673
1.096 | -87.0754
1.0997 | -86.9349
1.1034 | -86.9174
1.107 | -87.0584
1.1107 | -87.0969
1.1144 | -86.9768
1.118 | -86.9345
1.1217 | -87.0783
1.1254 | -87.2988
1.129 | -87.5293
1.1327 | -87.854
1.1364 | -88.1957
1.14 | -88.3292
1.1437 | -88.1933
1.1474 | -87.9846
1.151 | -87.9207
1.1547 | -88.04
1.1584 | -88.1793
1.162 | -88.095
1.1657 | -87.7877
1.1694 | -87.5327
1.173 | -87.5283
1.1767 | -87.772
1.1804 | -88.0856
1.184 | -88.3029
1.1877 | -88.3434
1.1913 | -88.1039
1.195 | -87.6604
1.1987 | -87.3284
1.2023 | -87.3692
1.206 | -87.7393
1.2097 | -88.1177
1.2133 | -88.2087
1.217 | -88.019
1.2207 | -87.8302
1.2243 | -87.8655
1.228 | -87.9784
1.2317 | -87.8779
1.2353 | -87.6102
1.239 | -87.4618
1.2427 | -87.6263
1.2463 | -88.0371
1.25 | -88.4207
1.2537 | -88.6007
1.2573 | -88.5581
1.261 | -88.3666
1.2647 | -88.2966
1.2683 | -88.5435
1.272 | -88.9706
1.2757 | -89.2933
1.2793 | -89.3379
1.283 | -89.0955
1.2867 | -88.7826
1.2903 | -88.552
1.294 | -88.3765
1.2977 | -88.2313
1.3013 | -88.211
1.305 | -88.4007
1.3087 | -88.7437
1.3123 | -89.0665
1.316 | -89.1845
1.3196 | -89.0448
1.3233 | -88.8894
1.327 | -88.9139
1.3306 | -89.0204
1.3343 | -89.1353
1.338 | -89.4411
1.3416 | -89.9662
1.3453 | -90.2965
1.349 | -90.0946
1.3526 | -89.6726
1.3563 | -89.434
1.36 | -89.4211
1.3636 | -89.4704
1.3673 | -89.4456
1.371 | -89.3326
1.3746 | -89.2588
1.3783 | -89.314
1.382 | -89.411
1.3856 | -89.4774
1.3893 | -89.5793
1.393 | -89.7654
1.3966 | -90.0216
1.4003 | -90.1916
1.404 | -90.1166
1.4076 | -89.8627
1.4113 | -89.5474
1.415 | -89.3166
1.4186 | -89.2877
1.4223 | -89.3375
1.426 | -89.468
1.4296 | -89.7664
1.4333 | -89.9648
1.4369 | -89.8012
1.4406 | -89.4995
1.4443 | -89.3507
1.4479 | -89.3546
1.4516 | -89.3122
1.4553 | -89.1478
1.4589 | -89.0444
1.4626 | -89.2317
1.4663 | -89.6427
1.4699 | -89.9391
1.4736 | -89.9245
1.4773 | -89.8279
1.4809 | -89.9193
1.4846 | -90.108
1.4883 | -90.1276
1.4919 | -90.0069
1.4956 | -90.0305
1.4993 | -90.3108
1.5029 | -90.6006
1.5066 | -90.6637
1.5103 | -90.5782
1.5139 | -90.5518
1.5176 | -90.6402
1.5213 | -90.7167
1.5249 | -90.6009
1.5286 | -90.311
1.5323 | -90.0773
1.5359 | -90.0302
1.5396 | -90.1073
1.5433 | -90.3179
1.5469 | -90.77
1.5506 | -91.3651
1.5543 | -91.8339
1.5579 | -92.0526
1.5616 | -91.9393
1.5652 | -91.5964
1.5689 | -91.3399
1.5726 | -91.2658
1.5762 | -91.2582
1.5799 | -91.1688
1.5836 | -91.111
1.5872 | -91.1337
1.5909 | -91.1879
1.5946 | -91.1677
1.5982 | -91.0162
1.6019 | -90.7963
1.6056 | -90.56
1.6092 | -90.3488
1.6129 | -90.2079
1.6166 | -90.1683
1.6202 | -90.2605
1.6239 | -90.473
1.6276 | -90.7157
1.6312 | -90.9331
1.6349 | -91.162
1.6386 | -91.4518
1.6422 | -91.7258
1.6459 | -91.8532
1.6496 | -91.7774
1.6532 | -91.5414
1.6569 | -91.2653
1.6606 | -91.0646
1.6642 | -90.9726
1.6679 | -90.9811
1.6716 | -91.0832
1.6752 | -91.2174
1.6789 | -91.3023
1.6826 | -91.3184
1.6862 | -91.3355
1.6899 | -91.3991
1.6935 | -91.4634
1.6972 | -91.4559
1.7009 | -91.3716
1.7045 | -91.2755
1.7082 | -91.2541
1.7119 | -91.2827
1.7155 | -91.2667
1.7192 | -91.1462
1.7229 | -90.996
1.7265 | -90.9284
1.7302 | -90.9772
1.7339 | -91.077
1.7375 | -91.1912
1.7412 | -91.3145
1.7449 | -91.4371
1.7485 | -91.5757
1.7522 | -91.7666
1.7559 | -91.9757
1.7595 | -92.0772
1.7632 | -91.9548
1.7669 | -91.666
1.7705 | -91.3619
1.7742 | -91.1523
1.7779 | -91.0792
1.7815 | -91.1202
1.7852 | -91.2166
1.7889 | -91.2942
1.7925 | -91.3424
1.7962 | -91.3977
1.7999 | -91.4849
1.8035 | -91.6418
1.8072 | -91.9094
1.8109 | -92.2309
1.8145 | -92.4512
1.8182 | -92.464
1.8218 | -92.3501
1.8255 | -92.2842
1.8292 | -92.3836
1.8328 | -92.622
1.8365 | -92.8594
1.8402 | -92.9707
1.8438 | -92.9115
1.8475 | -92.7147
1.8512 | -92.4649
1.8548 | -92.2463
1.8585 | -92.129
1.8622 | -92.1424
1.8658 | -92.2287
1.8695 | -92.3423
1.8732 | -92.4697
1.8768 | -92.5861
1.8805 | -92.6132
1.8842 | -92.4972
1.8878 | -92.3283
1.8915 | -92.2275
1.8952 | -92.2155
1.8988 | -92.2551
1.9025 | -92.2738
1.9062 | -92.2311
1.9098 | -92.0964
1.9135 | -91.934
1.9172 | -91.8851
1.9208 | -92.0096
1.9245 | -92.2052
1.9282 | -92.2927
1.9318 | -92.2027
1.9355 | -92.006
1.9391 | -91.8302
1.9428 | -91.7825
1.9465 | -91.9015
1.9501 | -92.1629
1.9538 | -92.4684
1.9575 | -92.6812
1.9611 | -92.7267
1.9648 | -92.6853
1.9685 | -92.6999
1.9721 | -92.8105
1.9758 | -92.9332
1.9795 | -92.9478
1.9831 | -92.8113
1.9868 | -92.6023
1.9905 | -92.4377
1.9941 | -92.3794
1.9978 | -92.385
2.0015 | -92.4192
2.0051 | -92.5153
2.0088 | -92.7546
2.0125 | -93.1285
2.0161 | -93.5273
2.0198 | -93.8167
2.0235 | -93.9058
2.0271 | -93.7755
2.0308 | -93.5318
2.0345 | -93.2981
2.0381 | -93.1535
2.0418 | -93.0715
2.0455 | -93.0106
2.0491 | -92.975
2.0528 | -92.9647
2.0565 | -92.926
2.0601 | -92.7893
2.0638 | -92.5808
2.0674 | -92.415
2.0711 | -92.361
2.0748 | -92.4196
2.0784 | -92.571
2.0821 | -92.7924
2.0858 | -93.0027
2.0894 | -93.029
2.0931 | -92.8237
2.0968 | -92.5179
2.1004 | -92.3083
2.1041 | -92.3183
2.1078 | -92.5449
2.1114 | -92.851
2.1151 | -92.9796
2.1188 | -93.1138
2.1224 | -93.2105
2.1261 | -93.3024
2.1298 | -93.3602
2.1334 | -93.3613
2.1371 | -93.2859
2.1408 | -93.1574
2.1444 | -93.0494
2.1481 | -93.0276
2.1518 | -93.1201
2.1554 | -93.2965
2.1591 | -93.4686
2.1628 | -93.5401
2.1664 | -93.4849
2.1701 | -93.3512
2.1738 | -93.189
2.1774 | -93.0559
2.1811 | -92.969
2.1848 | -92.933
2.1884 | -92.9898
2.1921 | -93.1681
2.1957 | -93.4524
2.1994 | -93.7847
2.2031 | -94.0844
2.2067 | -94.2898
2.2104 | -94.3688
2.2141 | -94.3302
2.2177 | -94.208
2.2214 | -94.0667
2.2251 | -93.946
2.2287 | -93.8217
2.2324 | -93.6928
2.2361 | -93.6006
2.2397 | -93.5883
2.2434 | -93.6659
2.2471 | -93.8015
2.2507 | -93.9329
2.2544 | -93.983
2.2581 | -93.8965
2.2617 | -93.7148
2.2654 | -93.5262
2.2691 | -93.4269
2.2727 | -93.4524
2.2764 | -93.5795
2.2801 | -93.7465
2.2837 | -93.8878
2.2874 | -93.9341
2.2911 | -93.8862
2.2947 | -93.7874
2.2984 | -93.7075
2.3021 | -93.6897
2.3057 | -93.7499
2.3094 | -93.8634
2.3131 | -94.0063
2.3167 | -94.136
2.3204 | -94.2282
2.324 | -94.2559
2.3277 | -94.2284
2.3314 | -94.1538
2.335 | -94.0709
2.3387 | -94.0014
2.3424 | -93.9526
2.346 | -93.9132
2.3497 | -93.8779
2.3534 | -93.8651
2.357 | -93.9153
2.3607 | -94.0254
2.3644 | -94.1684
2.368 | -94.2966
2.3717 | -94.3762
2.3754 | -94.3793
2.379 | -94.2994
2.3827 | -94.1748
2.3864 | -94.0491
2.39 | -93.9474
2.3937 | -93.8783
2.3974 | -93.824
2.401 | -93.7794
2.4047 | -93.7608
2.4084 | -93.7856
2.412 | -93.8574
2.4157 | -93.9679
2.4194 | -94.1115
2.423 | -94.2648
2.4267 | -94.4094
2.4304 | -94.5215
2.434 | -94.5707
2.4377 | -94.5669
2.4413 | -94.5493
2.445 | -94.5516
2.4487 | -94.601
2.4523 | -94.6867
2.456 | -94.7808
2.4597 | -94.8246
2.4633 | -94.7846
2.467 | -94.6252
2.4707 | -94.3537
2.4743 | -94.0492
2.478 | -93.8099
2.4817 | -93.7109
2.4853 | -93.7842
2.489 | -93.9842
2.4927 | -94.2254
2.4963 | -94.4209
2.5 | -94.5402
2.5037 | -94.5978
2.5073 | -94.6386
2.511 | -94.6804
2.5147 | -94.728
2.5183 | -94.7559
2.522 | -94.7128
2.5257 | -94.6129
2.5293 | -94.5107
2.533 | -94.4544
2.5367 | -94.4591
2.5403 | -94.5123
2.544 | -94.579
2.5477 | -94.6365
2.5513 | -94.6748
2.555 | -94.6974
2.5587 | -94.7393
2.5623 | -94.7936
2.566 | -94.8337
2.5696 | -94.8548
2.5733 | -94.8844
2.577 | -94.9442
2.5806 | -95.0244
2.5843 | -95.1089
2.588 | -95.1994
2.5916 | -95.3238
2.5953 | -95.4431
2.599 | -95.4846
2.6026 | -95.4319
2.6063 | -95.2732
2.61 | -95.0573
2.6136 | -94.8662
2.6173 | -94.7583
2.621 | -94.737
2.6246 | -94.7854
2.6283 | -94.8633
2.632 | -94.9372
2.6356 | -94.9926
2.6393 | -95.0097
2.643 | -94.9969
2.6466 | -95.0014
2.6503 | -95.0099
2.654 | -95.0276
2.6576 | -95.0477
2.6613 | -95.0709
2.665 | -95.1016
2.6686 | -95.1285
2.6723 | -95.1562
2.676 | -95.1663
2.6796 | -95.1672
2.6833 | -95.1383
2.6869 | -95.0666
2.6906 | -94.9638
2.6943 | -94.8641
2.6979 | -94.8058
2.7016 | -94.8134
2.7053 | -94.8775
2.7089 | -94.9852
2.7126 | -95.0911
2.7163 | -95.15
2.7199 | -95.136
2.7236 | -95.0799
2.7273 | -95.002
2.7309 | -94.9493
2.7346 | -94.9382
2.7383 | -94.9673
2.7419 | -95.0516
2.7456 | -95.1874
2.7493 | -95.3384
2.7529 | -95.4848
2.7566 | -95.5835
2.7603 | -95.6423
2.7639 | -95.6713
2.7676 | -95.6922
2.7713 | -95.7225
2.7749 | -95.7705
2.7786 | -95.8087
2.7823 | -95.8124
2.7859 | -95.7828
2.7896 | -95.7236
2.7933 | -95.6526
2.7969 | -95.5903
2.8006 | -95.5185
2.8043 | -95.4497
2.8079 | -95.3931
2.8116 | -95.3711
2.8152 | -95.3866
2.8189 | -95.4197
2.8226 | -95.4616
2.8262 | -95.5016
2.8299 | -95.5245
2.8336 | -95.5145
2.8372 | -95.4743
2.8409 | -95.4074
2.8446 | -95.3137
2.8482 | -95.2197
2.8519 | -95.1428
2.8556 | -95.0996
2.8592 | -95.1072
2.8629 | -95.1638
2.8666 | -95.2461
2.8702 | -95.3274
2.8739 | -95.3879
2.8776 | -95.415
2.8812 | -95.434
2.8849 | -95.4664
2.8886 | -95.5238
2.8922 | -95.5957
2.8959 | -95.6545
2.8996 | -95.692
2.9032 | -95.7105
2.9069 | -95.7367
2.9106 | -95.8081
2.9142 | -95.9403
2.9179 | -96.1376
2.9216 | -96.3804
2.9252 | -96.6089
2.9289 | -96.7374
2.9326 | -96.7199
2.9362 | -96.559
2.9399 | -96.3219
2.9435 | -96.0913
2.9472 | -95.9212
2.9509 | -95.8482
2.9545 | -95.8454
2.9582 | -95.8731
2.9619 | -95.9088
2.9655 | -95.9338
2.9692 | -95.9531
2.9729 | -95.9794
2.9765 | -96.0068
2.9802 | -96.0376
2.9839 | -96.0537
2.9875 | -96.0174
2.9912 | -95.9245
2.9949 | -95.7942
2.9985 | -95.6764
3.0022 | -95.6153
3.0059 | -95.6127
3.0095 | -95.6598
3.0132 | -95.7414
3.0169 | -95.8346
3.0205 | -95.9182
3.0242 | -95.9749
3.0279 | -96.003
3.0315 | -96.0198
3.0352 | -96.0572
3.0389 | -96.1251
3.0425 | -96.2392
3.0462 | -96.3777
3.0499 | -96.5105
3.0535 | -96.609
3.0572 | -96.6502
3.0608 | -96.6262
3.0645 | -96.5786
3.0682 | -96.5111
3.0718 | -96.4112
3.0755 | -96.2983
3.0792 | -96.1961
3.0828 | -96.1138
3.0865 | -96.0568
3.0902 | -96.0224
3.0938 | -96.0029
3.0975 | -96.0039
3.1012 | -96.0174
3.1048 | -96.0419
3.1085 | -96.076
3.1122 | -96.1
3.1158 | -96.1378
3.1195 | -96.2018
3.1232 | -96.2725
3.1268 | -96.3464
3.1305 | -96.408
3.1342 | -96.4568
3.1378 | -96.4822
3.1415 | -96.4762
3.1452 | -96.4306
3.1488 | -96.3716
3.1525 | -96.3083
3.1562 | -96.276
3.1598 | -96.2767
3.1635 | -96.2911
3.1672 | -96.3312
3.1708 | -96.4492
3.1745 | -96.5526
3.1782 | -96.6793
3.1818 | -96.8102
3.1855 | -96.9087
3.1892 | -96.9694
3.1928 | -96.9806
3.1965 | -96.9484
3.2001 | -96.8715
3.2038 | -96.772
3.2075 | -96.6617
3.2111 | -96.5658
3.2148 | -96.4938
3.2185 | -96.4433
3.2221 | -96.4389
3.2258 | -96.4786
3.2295 | -96.5654
3.2331 | -96.6881
3.2368 | -96.8416
3.2405 | -97.001
3.2441 | -97.1275
3.2478 | -97.2139
3.2515 | -97.2362
3.2551 | -97.2011
3.2588 | -97.1304
3.2625 | -97.0336
3.2661 | -96.9346
3.2698 | -96.8573
3.2735 | -96.8152
3.2771 | -96.7954
3.2808 | -96.8043
3.2845 | -96.8378
3.2881 | -96.8732
3.2918 | -96.8961
3.2955 | -96.8992
3.2991 | -96.8799
3.3028 | -96.839
3.3065 | -96.7938
3.3101 | -96.7476
3.3138 | -96.7167
3.3174 | -96.7012
3.3211 | -96.702
3.3248 | -96.7156
3.3284 | -96.749
3.3321 | -96.8
3.3358 | -96.8602
3.3394 | -96.922
3.3431 | -96.9785
3.3468 | -97.0264
3.3504 | -97.0705
3.3541 | -97.1033
3.3578 | -97.1188
3.3614 | -97.1174
3.3651 | -97.1134
3.3688 | -97.1043
3.3724 | -97.1
3.3761 | -97.1037
3.3798 | -97.1106
3.3834 | -97.104
3.3871 | -97.0658
3.3908 | -96.995
3.3944 | -96.8997
3.3981 | -96.8093
3.4018 | -96.741
3.4054 | -96.7059
3.4091 | -96.7241
3.4128 | -96.7848
3.4164 | -96.8609
3.4201 | -96.9456
3.4238 | -97.0341
3.4274 | -97.1252
3.4311 | -97.1995
3.4348 | -97.2605
3.4384 | -97.3196
3.4421 | -97.3813
3.4457 | -97.4494
3.4494 | -97.5368
3.4531 | -97.6428
3.4567 | -97.7664
3.4604 | -97.8989
3.4641 | -98.0008
3.4677 | -98.0753
3.4714 | -98.1141
3.4751 | -98.0778
3.4787 | -97.9897
3.4824 | -97.8734
3.4861 | -97.7503
3.4897 | -97.6286
3.4934 | -97.5166
3.4971 | -97.4364
3.5007 | -97.3922
3.5044 | -97.3867
3.5081 | -97.4285
3.5117 | -97.5047
3.5154 | -97.5819
3.5191 | -97.6515
3.5227 | -97.6868
3.5264 | -97.7004
3.5301 | -97.694
3.5337 | -97.674
3.5374 | -97.639
3.5411 | -97.5911
3.5447 | -97.5365
3.5484 | -97.4753
3.5521 | -97.4248
3.5557 | -97.4003
3.5594 | -97.3985
3.5631 | -97.3984
3.5667 | -97.3978
3.5704 | -97.4005
3.574 | -97.3914
3.5777 | -97.3844
3.5814 | -97.3893
3.585 | -97.4233
3.5887 | -97.4788
3.5924 | -97.5514
3.596 | -97.6244
3.5997 | -97.6688
3.6034 | -97.6875
3.607 | -97.667
3.6107 | -97.6183
3.6144 | -97.561
3.618 | -97.5054
3.6217 | -97.4632
3.6254 | -97.4346
3.629 | -97.4269
3.6327 | -97.436
3.6364 | -97.4522
3.64 | -97.4742
3.6437 | -97.4774
3.6474 | -97.4731
3.651 | -97.4587
3.6547 | -97.4418
3.6584 | -97.4286
3.662 | -97.4443
3.6657 | -97.5004
3.6694 | -97.5962
3.673 | -97.7127
3.6767 | -97.8211
3.6804 | -97.9077
3.684 | -97.9478
3.6877 | -97.9199
3.6913 | -97.8412
3.695 | -97.7362
3.6987 | -97.6565
3.7023 | -97.5993
3.706 | -97.5484
3.7097 | -97.5141
3.7133 | -97.5002
3.717 | -97.4965
3.7207 | -97.4894
3.7243 | -97.4864
3.728 | -97.5031
3.7317 | -97.5471
3.7353 | -97.6282
3.739 | -97.7473
3.7427 | -97.9001
3.7463 | -98.0765
3.75 | -98.2533
3.7537 | -98.3992
3.7573 | -98.5027
3.761 | -98.5519
3.7647 | -98.5421
3.7683 | -98.486
3.772 | -98.3874
3.7757 | -98.2745
3.7793 | -98.165
3.783 | -98.0786
3.7867 | -98.0153
3.7903 | -97.9776
3.794 | -97.9648
3.7977 | -97.9594
3.8013 | -97.957
3.805 | -97.9561
3.8087 | -97.9524
3.8123 | -97.9422
3.816 | -97.9352
3.8196 | -97.9373
3.8233 | -97.9358
3.827 | -97.9408
3.8306 | -97.9456
3.8343 | -97.9359
3.838 | -97.9054
3.8416 | -97.8623
3.8453 | -97.8185
3.849 | -97.7797
3.8526 | -97.75
3.8563 | -97.7397
3.86 | -97.753
3.8636 | -97.785
3.8673 | -97.825
3.871 | -97.8744
3.8746 | -97.9293
3.8783 | -97.9909
3.882 | -98.0521
3.8856 | -98.1195
3.8893 | -98.183
3.893 | -98.2355
3.8966 | -98.2742
3.9003 | -98.2951
3.904 | -98.2937
3.9076 | -98.2682
3.9113 | -98.2281
3.915 | -98.1803
3.9186 | -98.1236
3.9223 | -98.0616
3.926 | -98.005
3.9296 | -97.9615
3.9333 | -97.9309
3.9369 | -97.9157
3.9406 | -97.9258
3.9443 | -97.9652
3.9479 | -98.0368
3.9516 | -98.1367
3.9553 | -98.2613
3.9589 | -98.3898
3.9626 | -98.5012
3.9663 | -98.5883
3.9699 | -98.6411
3.9736 | -98.663
3.9773 | -98.6593
3.9809 | -98.6409
3.9846 | -98.6116
3.9883 | -98.5752
3.9919 | -98.5465
3.9956 | -98.5291
3.9993 | -98.5172
4.0029 | -98.5114
4.0066 | -98.5103
4.0103 | -98.5005
4.0139 | -98.4837
4.0176 | -98.4546
4.0213 | -98.4049
4.0249 | -98.3415
4.0286 | -98.2773
4.0323 | -98.2211
4.0359 | -98.1772
4.0396 | -98.1555
4.0433 | -98.1636
4.0469 | -98.1975
4.0506 | -98.2525
4.0543 | -98.3168
4.0579 | -98.3941
4.0616 | -98.4759
4.0652 | -98.5354
4.0689 | -98.5705
4.0726 | -98.5878
4.0762 | -98.589
4.0799 | -98.5914
4.0836 | -98.6164
4.0872 | -98.668
4.0909 | -98.7486
4.0946 | -98.8463
4.0982 | -98.9506
4.1019 | -99.0494
4.1056 | -99.1139
4.1092 | -99.146
4.1129 | -99.1465
4.1166 | -99.1203
4.1202 | -99.0866
4.1239 | -99.0602
4.1276 | -99.0496
4.1312 | -99.0544
4.1349 | -99.0594
4.1386 | -99.0609
4.1422 | -99.0471
4.1459 | -99.0219
4.1496 | -98.9989
4.1532 | -98.9775
4.1569 | -98.9701
4.1606 | -98.9766
4.1642 | -98.979
4.1679 | -98.9799
4.1716 | -98.9748
4.1752 | -98.9466
4.1789 | -98.9032
4.1826 | -98.8471
4.1862 | -98.7827
4.1899 | -98.7157
4.1935 | -98.6614
4.1972 | -98.6227
4.2009 | -98.5957
4.2045 | -98.5771
4.2082 | -98.5607
4.2119 | -98.5483
4.2155 | -98.5527
4.2192 | -98.5711
4.2229 | -98.5983
4.2265 | -98.6408
4.2302 | -98.6962
4.2339 | -98.7473
4.2375 | -98.79
4.2412 | -98.8351
4.2449 | -98.8747
4.2485 | -98.9052
4.2522 | -98.9351
4.2559 | -98.9566
4.2595 | -98.9584
4.2632 | -98.9535
4.2669 | -98.9399
4.2705 | -98.9115
4.2742 | -98.8755
4.2779 | -98.8326
4.2815 | -98.7938
4.2852 | -98.7756
4.2889 | -98.7819
4.2925 | -98.8163
4.2962 | -98.8698
4.2999 | -98.9446
4.3035 | -99.0342
4.3072 | -99.1305
4.3109 | -99.2256
4.3145 | -99.306
4.3182 | -99.3688
4.3218 | -99.4081
4.3255 | -99.4271
4.3292 | -99.4335
4.3328 | -99.4301
4.3365 | -99.434
4.3402 | -99.4597
4.3438 | -99.4992
4.3475 | -99.5539
4.3512 | -99.6044
4.3548 | -99.6389
4.3585 | -99.6536
4.3622 | -99.6317
4.3658 | -99.5795
4.3695 | -99.5072
4.3732 | -99.4304
4.3768 | -99.3634
4.3805 | -99.3051
4.3842 | -99.2705
4.3878 | -99.2512
4.3915 | -99.2388
4.3952 | -99.2331
4.3988 | -99.2305
4.4025 | -99.2327
4.4062 | -99.2356
4.4098 | -99.2399
4.4135 | -99.2563
4.4172 | -99.2886
4.4208 | -99.342
4.4245 | -99.4029
4.4282 | -99.4572
4.4318 | -99.5073
4.4355 | -99.5466
4.4391 | -99.5724
4.4428 | -99.5879
4.4465 | -99.5998
4.4501 | -99.6159
4.4538 | -99.6229
4.4575 | -99.6223
4.4611 | -99.6292
4.4648 | -99.6335
4.4685 | -99.6301
4.4721 | -99.6178
4.4758 | -99.5949
4.4795 | -99.5637
4.4831 | -99.5153
4.4868 | -99.4536
4.4905 | -99.3863
4.4941 | -99.3259
4.4978 | -99.2786
4.5015 | -99.2432
4.5051 | -99.2183
4.5088 | -99.2022
4.5125 | -99.19
4.5161 | -99.1839
4.5198 | -99.1943
4.5235 | -99.216
4.5271 | -99.2441
4.5308 | -99.2714
4.5345 | -99.2987
4.5381 | -99.3248
4.5418 | -99.3554
4.5455 | -99.3891
4.5491 | -99.4112
4.5528 | -99.4319
4.5565 | -99.4541
4.5601 | -99.4712
4.5638 | -99.4799
4.5674 | -99.4744
4.5711 | -99.4733
4.5748 | -99.4786
4.5784 | -99.4903
4.5821 | -99.5052
4.5858 | -99.518
4.5894 | -99.5294
4.5931 | -99.5509
4.5968 | -99.5799
4.6004 | -99.6103
4.6041 | -99.6492
4.6078 | -99.6924
4.6114 | -99.7255
4.6151 | -99.7434
4.6188 | -99.7471
4.6224 | -99.728
4.6261 | -99.688
4.6298 | -99.6373
4.6334 | -99.5839
4.6371 | -99.5294
4.6408 | -99.4748
4.6444 | -99.4335
4.6481 | -99.3966
4.6518 | -99.3664
4.6554 | -99.328
4.6591 | -99.2778
4.6628 | -99.2256
4.6664 | -99.1727
4.6701 | -99.1325
4.6738 | -99.1077
4.6774 | -99.1002
4.6811 | -99.1038
4.6848 | -99.1131
4.6884 | -99.1162
4.6921 | -99.1177
4.6957 | -99.1216
4.6994 | -99.1275
4.7031 | -99.1402
4.7067 | -99.1664
4.7104 | -99.2044
4.7141 | -99.2578
4.7177 | -99.3255
4.7214 | -99.395
4.7251 | -99.4549
4.7287 | -99.4912
4.7324 | -99.5029
4.7361 | -99.4929
4.7397 | -99.4636
4.7434 | -99.43
4.7471 | -99.4023
4.7507 | -99.3848
4.7544 | -99.3796
4.7581 | -99.4114
4.7617 | -99.4341
4.7654 | -99.4668
4.7691 | -99.5052
4.7727 | -99.5543
4.7764 | -99.61
4.7801 | -99.6623
4.7837 | -99.6971
4.7874 | -99.7197
4.7911 | -99.7225
4.7947 | -99.7117
4.7984 | -99.7026
4.8021 | -99.7017
4.8057 | -99.7094
4.8094 | -99.7348
4.813 | -99.7756
4.8167 | -99.8234
4.8204 | -99.8721
4.824 | -99.9192
4.8277 | -99.9607
4.8314 | -99.9943
4.835 | -100.0158
4.8387 | -100.0178
4.8424 | -100.0076
4.846 | -100.0016
4.8497 | -99.9901
4.8534 | -99.9765
4.857 | -99.9658
4.8607 | -99.9531
4.8644 | -99.9377
4.868 | -99.9313
4.8717 | -99.9339
4.8754 | -99.943
4.879 | -99.9618
4.8827 | -99.9951
4.8864 | -100.0328
4.89 | -100.0729
4.8937 | -100.1176
4.8974 | -100.1567
4.901 | -100.187
4.9047 | -100.2026
4.9084 | -100.2072
4.912 | -100.1943
4.9157 | -100.1643
4.9194 | -100.1262
4.923 | -100.0786
4.9267 | -100.0253
4.9304 | -99.9684
4.934 | -99.9085
4.9377 | -99.8574
4.9413 | -99.8109
4.945 | -99.7691
4.9487 | -99.7311
4.9523 | -99.7044
4.956 | -99.6968
4.9597 | -99.7087
4.9633 | -99.7373
4.967 | -99.7766
4.9707 | -99.8229
4.9743 | -99.8798
4.978 | -99.9394
4.9817 | -100.0043
4.9853 | -100.0682
4.989 | -100.1388
4.9927 | -100.2135
4.9963 | -100.2754
5 | -100.3228
5.0037 | -100.3579
5.0073 | -100.376
5.011 | -100.3748
5.0147 | -100.3579
5.0183 | -100.3318
5.022 | -100.3092
5.0257 | -100.2892
5.0293 | -100.2609
5.033 | -100.233
5.0367 | -100.1976
5.0403 | -100.1566
5.044 | -100.1112
5.0477 | -100.059
5.0513 | -100.0004
5.055 | -99.9437
5.0587 | -99.8897
5.0623 | -99.8507
5.066 | -99.8205
5.0696 | -99.7961
5.0733 | -99.7942
5.077 | -99.8083
5.0806 | -99.8271
5.0843 | -99.8435
5.088 | -99.8607
5.0916 | -99.8783
5.0953 | -99.8963
5.099 | -99.9119
5.1026 | -99.9358
5.1063 | -99.9804
5.11 | -100.0324
5.1136 | -100.1008
5.1173 | -100.1758
5.121 | -100.2398
5.1246 | -100.2921
5.1283 | -100.3335
5.132 | -100.3491
5.1356 | -100.3381
5.1393 | -100.307
5.143 | -100.2742
5.1466 | -100.2473
5.1503 | -100.214
5.154 | -100.1852
5.1576 | -100.1684
5.1613 | -100.1566
5.165 | -100.1418
5.1686 | -100.131
5.1723 | -100.1172
5.176 | -100.1074
5.1796 | -100.1054
5.1833 | -100.1065
5.187 | -100.121
5.1906 | -100.1477
5.1943 | -100.1908
5.1979 | -100.2501
5.2016 | -100.3142
5.2053 | -100.382
5.2089 | -100.4576
5.2126 | -100.5395
5.2163 | -100.6155
5.2199 | -100.6813
5.2236 | -100.7366
5.2273 | -100.7875
5.2309 | -100.8284
5.2346 | -100.852
5.2383 | -100.8614
5.2419 | -100.8522
5.2456 | -100.8178
5.2493 | -100.7807
5.2529 | -100.7405
5.2566 | -100.6949
5.2603 | -100.6577
5.2639 | -100.6264
5.2676 | -100.614
5.2713 | -100.6213
5.2749 | -100.6351
5.2786 | -100.6547
5.2823 | -100.6833
5.2859 | -100.698
5.2896 | -100.7113
5.2933 | -100.7165
5.2969 | -100.7218
5.3006 | -100.725
5.3043 | -100.7348
5.3079 | -100.7512
5.3116 | -100.7776
5.3152 | -100.8115
5.3189 | -100.8493
5.3226 | -100.8977
5.3262 | -100.951
5.3299 | -101.0001
5.3336 | -101.0451
5.3372 | -101.0913
5.3409 | -101.1346
5.3446 | -101.1733
5.3482 | -101.2037
5.3519 | -101.2267
5.3556 | -101.2554
5.3592 | -101.2806
5.3629 | -101.293
5.3666 | -101.3044
5.3702 | -101.3115
5.3739 | -101.306
5.3776 | -101.2876
5.3812 | -101.2499
5.3849 | -101.2055
5.3886 | -101.1452
5.3922 | -101.0763
5.3959 | -101.0044
5.3996 | -100.9342
5.4032 | -100.8775
5.4069 | -100.8282
5.4106 | -100.7903
5.4142 | -100.7675
5.4179 | -100.7573
5.4216 | -100.7616
5.4252 | -100.7739
5.4289 | -100.7962
5.4326 | -100.8268
5.4362 | -100.8618
5.4399 | -100.8991
5.4435 | -100.9429
5.4472 | -100.9816
5.4509 | -101.0192
5.4545 | -101.0641
5.4582 | -101.104
5.4619 | -101.1391
5.4655 | -101.1746
5.4692 | -101.206
5.4729 | -101.2311
5.4765 | -101.2482
5.4802 | -101.255
5.4839 | -101.2499
5.4875 | -101.2291
5.4912 | -101.2062
5.4949 | -101.1819
5.4985 | -101.1526
5.5022 | -101.122
5.5059 | -101.0907
5.5095 | -101.0617
5.5132 | -101.0352
5.5169 | -101.0011
5.5205 | -100.9635
5.5242 | -100.918
5.5279 | -100.8719
5.5315 | -100.8262
5.5352 | -100.7728
5.5389 | -100.7107
5.5425 | -100.6529
5.5462 | -100.597
5.5499 | -100.5477
5.5535 | -100.5099
5.5572 | -100.4958
5.5609 | -100.4996
5.5645 | -100.5187
5.5682 | -100.5546
5.5718 | -100.6065
5.5755 | -100.6727
5.5792 | -100.755
5.5828 | -100.8533
5.5865 | -100.9593
5.5902 | -101.0617
5.5938 | -101.1588
5.5975 | -101.2433
5.6012 | -101.3203
5.6048 | -101.3718
5.6085 | -101.4008
5.6122 | -101.4058
5.6158 | -101.3987
5.6195 | -101.3849
5.6232 | -101.3611
5.6268 | -101.339
5.6305 | -101.3192
5.6342 | -101.2965
5.6378 | -101.279
5.6415 | -101.2656
5.6452 | -101.2634
5.6488 | -101.2707
5.6525 | -101.2773
5.6562 | -101.2846
5.6598 | -101.2797
5.6635 | -101.2604
5.6672 | -101.2383
5.6708 | -101.2147
5.6745 | -101.1878
5.6782 | -101.1571
5.6818 | -101.1273
5.6855 | -101.1081
5.6891 | -101.1018
5.6928 | -101.107
5.6965 | -101.1274
5.7001 | -101.1659
5.7038 | -101.2121
5.7075 | -101.2653
5.7111 | -101.3106
5.7148 | -101.3588
5.7185 | -101.3989
5.7221 | -101.4285
5.7258 | -101.4479
5.7295 | -101.4483
5.7331 | -101.4352
5.7368 | -101.4164
5.7405 | -101.3863
5.7441 | -101.3524
5.7478 | -101.3137
5.7515 | -101.2611
5.7551 | -101.2016
5.7588 | -101.1435
5.7625 | -101.0912
5.7661 | -101.0437
5.7698 | -100.9979
5.7735 | -100.9656
5.7771 | -100.9458
5.7808 | -100.9317
5.7845 | -100.9296
5.7881 | -100.94
5.7918 | -100.9617
5.7955 | -100.9849
5.7991 | -101.0012
5.8028 | -101.0277
5.8065 | -101.0496
5.8101 | -101.064
5.8138 | -101.0447
5.8174 | -101.0331
5.8211 | -101.013
5.8248 | -100.9867
5.8284 | -100.9507
5.8321 | -100.9058
5.8358 | -100.8502
5.8394 | -100.7865
5.8431 | -100.7233
5.8468 | -100.6661
5.8504 | -100.6084
5.8541 | -100.5566
5.8578 | -100.5109
5.8614 | -100.4734
5.8651 | -100.4477
5.8688 | -100.4336
5.8724 | -100.4284
5.8761 | -100.4394
5.8798 | -100.4509
5.8834 | -100.4733
5.8871 | -100.5011
5.8908 | -100.5322
5.8944 | -100.5701
5.8981 | -100.6064
5.9018 | -100.6554
5.9054 | -100.7108
5.9091 | -100.7667
5.9128 | -100.8201
5.9164 | -100.8703
5.9201 | -100.9173
5.9238 | -100.956
5.9274 | -100.9915
5.9311 | -101.0212
5.9348 | -101.0395
5.9384 | -101.0458
5.9421 | -101.0513
5.9457 | -101.0539
5.9494 | -101.0585
5.9531 | -101.061
5.9567 | -101.0582
5.9604 | -101.0562
5.9641 | -101.0491
5.9677 | -101.0331
5.9714 | -101.028
5.9751 | -101.0204
5.9787 | -101.017
5.9824 | -101.0116
5.9861 | -101.0103
5.9897 | -101.0273
5.9934 | -101.0452
5.9971 | -101.0667
6.0007 | -101.0974
6.0044 | -101.1355
6.0081 | -101.177
6.0117 | -101.2264
6.0154 | -101.2828
6.0191 | -101.3311
6.0227 | -101.3715
6.0264 | -101.4091
6.0301 | -101.443
6.0337 | -101.4666
6.0374 | -101.4867
6.0411 | -101.5026
6.0447 | -101.512
6.0484 | -101.5229
6.0521 | -101.5343
6.0557 | -101.5471
6.0594 | -101.561
6.063 | -101.5698
6.0667 | -101.5667
6.0704 | -101.5521
6.074 | -101.5421
6.0777 | -101.5297
6.0814 | -101.5084
6.085 | -101.487
6.0887 | -101.4729
6.0924 | -101.4642
6.096 | -101.4575
6.0997 | -101.4545
6.1034 | -101.4554
6.107 | -101.4619
6.1107 | -101.4736
6.1144 | -101.4892
6.118 | -101.5086
6.1217 | -101.5285
6.1254 | -101.5532
6.129 | -101.5822
6.1327 | -101.6081
6.1364 | -101.6358
6.14 | -101.652
6.1437 | -101.6669
6.1474 | -101.6795
6.151 | -101.6859
6.1547 | -101.6889
6.1584 | -101.6991
6.162 | -101.7092
6.1657 | -101.7138
6.1694 | -101.7256
6.173 | -101.7437
6.1767 | -101.7629
6.1804 | -101.7842
6.184 | -101.8122
6.1877 | -101.8398
6.1913 | -101.8599
6.195 | -101.868
6.1987 | -101.8683
6.2023 | -101.8644
6.206 | -101.8543
6.2097 | -101.8404
6.2133 | -101.8285
6.217 | -101.8101
6.2207 | -101.791
6.2243 | -101.7793
6.228 | -101.7708
6.2317 | -101.7781
6.2353 | -101.7866
6.239 | -101.8012
6.2427 | -101.822
6.2463 | -101.8459
6.25 | -101.8697
6.2537 | -101.8927
6.2573 | -101.9097
6.261 | -101.9132
6.2647 | -101.908
6.2683 | -101.9027
6.272 | -101.8882
6.2757 | -101.8637
6.2793 | -101.8434
6.283 | -101.8281
6.2867 | -101.8093
6.2903 | -101.7859
6.294 | -101.7648
6.2977 | -101.7468
6.3013 | -101.7251
6.305 | -101.71
6.3087 | -101.7011
6.3123 | -101.6897
6.316 | -101.6843
6.3196 | -101.687
6.3233 | -101.6916
6.327 | -101.7048
6.3306 | -101.7204
6.3343 | -101.739
6.338 | -101.7623
6.3416 | -101.7898
6.3453 | -101.8078
6.349 | -101.8268
6.3526 | -101.8437
6.3563 | -101.8566
6.36 | -101.8571
6.3636 | -101.8495
6.3673 | -101.8413
6.371 | -101.8288
6.3746 | -101.8104
6.3783 | -101.7917
6.382 | -101.7727
6.3856 | -101.7515
6.3893 | -101.7285
6.393 | -101.7109
6.3966 | -101.7066
6.4003 | -101.7037
6.404 | -101.7117
6.4076 | -101.7257
6.4113 | -101.7419
6.415 | -101.7642
6.4186 | -101.7913
6.4223 | -101.8247
6.426 | -101.8628
6.4296 | -101.9002
6.4333 | -101.9367
6.437 | -101.9709
6.4406 | -102.0083
6.4443 | -102.048
6.4479 | -102.0808
6.4516 | -102.1089
6.4553 | -102.1358
6.4589 | -102.1607
6.4626 | -102.1732
6.4663 | -102.1741
6.4699 | -102.1707
6.4736 | -102.1627
6.4773 | -102.1536
6.4809 | -102.1489
6.4846 | -102.1494
6.4883 | -102.1591
6.4919 | -102.169
6.4956 | -102.1843
6.4993 | -102.2141
6.5029 | -102.2498
6.5066 | -102.2926
6.5103 | -102.3357
6.5139 | -102.3714
6.5176 | -102.4009
6.5213 | -102.4213
6.5249 | -102.4298
6.5286 | -102.4299
6.5323 | -102.4236
6.5359 | -102.4145
6.5396 | -102.3932
6.5433 | -102.3596
6.5469 | -102.324
6.5506 | -102.2846
6.5543 | -102.247
6.5579 | -102.2065
6.5616 | -102.1656
6.5652 | -102.1293
6.5689 | -102.0988
6.5726 | -102.0784
6.5762 | -102.0658
6.5799 | -102.065
6.5836 | -102.081
6.5872 | -102.1124
6.5909 | -102.1625
6.5946 | -102.2232
6.5982 | -102.2933
6.6019 | -102.3732
6.6056 | -102.4553
6.6092 | -102.5433
6.6129 | -102.6295
6.6166 | -102.7055
6.6202 | -102.7675
6.6239 | -102.8103
6.6276 | -102.8283
6.6312 | -102.8237
6.6349 | -102.8045
6.6386 | -102.7673
6.6422 | -102.7186
6.6459 | -102.6696
6.6496 | -102.6181
6.6532 | -102.563
6.6569 | -102.5157
6.6606 | -102.4782
6.6642 | -102.4522
6.6679 | -102.4353
6.6716 | -102.4266
6.6752 | -102.4345
6.6789 | -102.4615
6.6826 | -102.4927
6.6862 | -102.5326
6.6899 | -102.587
6.6935 | -102.6455
6.6972 | -102.6988
6.7009 | -102.7433
6.7045 | -102.7734
6.7082 | -102.7851
6.7119 | -102.782
6.7155 | -102.7684
6.7192 | -102.7392
6.7229 | -102.6971
6.7265 | -102.6463
6.7302 | -102.585
6.7339 | -102.5221
6.7375 | -102.4681
6.7412 | -102.4148
6.7449 | -102.3581
6.7485 | -102.3113
6.7522 | -102.272
6.7559 | -102.2441
6.7595 | -102.2227
6.7632 | -102.2112
6.7669 | -102.2096
6.7705 | -102.2188
6.7742 | -102.2364
6.7779 | -102.2536
6.7815 | -102.2761
6.7852 | -102.3068
6.7889 | -102.3365
6.7925 | -102.3645
6.7962 | -102.3889
6.7999 | -102.4164
6.8035 | -102.4384
6.8072 | -102.4514
6.8109 | -102.4689
6.8145 | -102.4764
6.8182 | -102.4818
6.8218 | -102.4883
6.8255 | -102.4857
6.8292 | -102.4807
6.8328 | -102.4712
6.8365 | -102.4551
6.8402 | -102.4361
6.8438 | -102.4177
6.8475 | -102.3963
6.8512 | -102.3775
6.8548 | -102.3628
6.8585 | -102.3496
6.8622 | -102.3304
6.8658 | -102.3185
6.8695 | -102.3133
6.8732 | -102.3374
6.8768 | -102.3503
6.8805 | -102.3737
6.8842 | -102.4068
6.8878 | -102.4427
6.8915 | -102.4891
6.8952 | -102.5473
6.8988 | -102.6111
6.9025 | -102.666
6.9062 | -102.7172
6.9098 | -102.7582
6.9135 | -102.7755
6.9172 | -102.7842
6.9208 | -102.7739
6.9245 | -102.744
6.9282 | -102.7023
6.9318 | -102.6462
6.9355 | -102.5801
6.9391 | -102.5143
6.9428 | -102.4451
6.9465 | -102.3713
6.9501 | -102.3017
6.9538 | -102.2436
6.9575 | -102.1943
6.9611 | -102.1542
6.9648 | -102.1281
6.9685 | -102.1044
6.9721 | -102.089
6.9758 | -102.0821
6.9795 | -102.0815
6.9831 | -102.0886
6.9868 | -102.0975
6.9905 | -102.1118
6.9941 | -102.1302
6.9978 | -102.1492
7.0015 | -102.1732
7.0051 | -102.201
7.0088 | -102.2318
7.0125 | -102.2665
7.0161 | -102.2996
7.0198 | -102.3341
7.0235 | -102.3623
7.0271 | -102.39
7.0308 | -102.4148
7.0345 | -102.4347
7.0381 | -102.4536
7.0418 | -102.4739
7.0455 | -102.4883
7.0491 | -102.5013
7.0528 | -102.5113
7.0565 | -102.5158
7.0601 | -102.5133
7.0638 | -102.501
7.0674 | -102.4897
7.0711 | -102.4691
7.0748 | -102.4475
7.0784 | -102.4267
7.0821 | -102.4029
7.0858 | -102.3746
7.0894 | -102.3511
7.0931 | -102.3395
7.0968 | -102.3261
7.1004 | -102.3172
7.1041 | -102.3176
7.1078 | -102.326
7.1114 | -102.3426
7.1151 | -102.3678
7.1188 | -102.3985
7.1224 | -102.4363
7.1261 | -102.4769
7.1298 | -102.5174
7.1334 | -102.5566
7.1371 | -102.5943
7.1408 | -102.63
7.1444 | -102.6612
7.1481 | -102.6825
7.1518 | -102.6971
7.1554 | -102.7074
7.1591 | -102.7148
7.1628 | -102.7118
7.1664 | -102.6984
7.1701 | -102.6853
7.1738 | -102.6688
7.1774 | -102.6574
7.1811 | -102.6412
7.1848 | -102.6249
7.1884 | -102.618
7.1921 | -102.6159
7.1957 | -102.6097
7.1994 | -102.6026
7.2031 | -102.6038
7.2067 | -102.6069
7.2104 | -102.6124
7.2141 | -102.6188
7.2177 | -102.6325
7.2214 | -102.6535
7.2251 | -102.6804
7.2287 | -102.7135
7.2324 | -102.7523
7.2361 | -102.7966
7.2397 | -102.8397
7.2434 | -102.8834
7.2471 | -102.9248
7.2507 | -102.9607
7.2544 | -102.9999
7.2581 | -103.0415
7.2617 | -103.0711
7.2654 | -103.0976
7.2691 | -103.1288
7.2727 | -103.16
7.2764 | -103.1844
7.2801 | -103.206
7.2837 | -103.2236
7.2874 | -103.2344
7.2911 | -103.2342
7.2947 | -103.2299
7.2984 | -103.2231
7.3021 | -103.2154
7.3057 | -103.198
7.3094 | -103.1716
7.313 | -103.1403
7.3167 | -103.1098
7.3204 | -103.0793
7.324 | -103.047
7.3277 | -103.0066
7.3314 | -102.9656
7.335 | -102.9323
7.3387 | -102.9071
7.3424 | -102.8918
7.346 | -102.883
7.3497 | -102.8847
7.3534 | -102.8916
7.357 | -102.9114
7.3607 | -102.943
7.3644 | -102.9819
7.368 | -103.028
7.3717 | -103.0668
7.3754 | -103.0982
7.379 | -103.1217
7.3827 | -103.138
7.3864 | -103.1418
7.39 | -103.1375
7.3937 | -103.1264
7.3974 | -103.1076
7.401 | -103.0786
7.4047 | -103.059
7.4084 | -103.039
7.412 | -103.0219
7.4157 | -103.0135
7.4194 | -103.0082
7.423 | -103.0068
7.4267 | -103.0086
7.4304 | -103.0157
7.434 | -103.0228
7.4377 | -103.0355
7.4413 | -103.053
7.445 | -103.0628
7.4487 | -103.0658
7.4523 | -103.0669
7.456 | -103.0688
7.4597 | -103.0572
7.4633 | -103.0481
7.467 | -103.0317
7.4707 | -103.0064
7.4743 | -102.9791
7.478 | -102.9533
7.4817 | -102.9279
7.4853 | -102.903
7.489 | -102.8809
7.4927 | -102.8558
7.4963 | -102.8325
7.5 | -102.8112
7.5037 | -102.7987
7.5073 | -102.7932
7.511 | -102.7916
7.5147 | -102.7959
7.5183 | -102.8034
7.522 | -102.8132
7.5257 | -102.8233
7.5293 | -102.8344
7.533 | -102.8508
7.5367 | -102.8617
7.5403 | -102.8678
7.544 | -102.8786
7.5477 | -102.8922
7.5513 | -102.9071
7.555 | -102.9124
7.5587 | -102.9207
7.5623 | -102.9337
7.566 | -102.9457
7.5696 | -102.9561
7.5733 | -102.9667
7.577 | -102.9707
7.5806 | -102.9737
7.5843 | -102.9741
7.588 | -102.9729
7.5916 | -102.9692
7.5953 | -102.9701
7.599 | -102.9739
7.6026 | -102.978
7.6063 | -102.9883
7.61 | -103.0036
7.6136 | -103.0215
7.6173 | -103.0434
7.621 | -103.0638
7.6246 | -103.0808
7.6283 | -103.0923
7.632 | -103.0974
7.6356 | -103.1035
7.6393 | -103.1055
7.643 | -103.1003
7.6466 | -103.0918
7.6503 | -103.0848
7.654 | -103.0756
7.6576 | -103.0569
7.6613 | -103.0411
7.665 | -103.0271
7.6686 | -103.0214
7.6723 | -103.0198
7.676 | -103.0256
7.6796 | -103.0382
7.6833 | -103.0626
7.687 | -103.099
7.6906 | -103.1448
7.6943 | -103.1952
7.6979 | -103.2525
7.7016 | -103.3149
7.7053 | -103.3832
7.7089 | -103.4436
7.7126 | -103.4947
7.7163 | -103.5382
7.7199 | -103.5701
7.7236 | -103.5873
7.7273 | -103.5882
7.7309 | -103.5732
7.7346 | -103.5405
7.7383 | -103.5065
7.7419 | -103.4723
7.7456 | -103.4339
7.7493 | -103.3838
7.7529 | -103.3341
7.7566 | -103.2905
7.7603 | -103.2515
7.7639 | -103.2201
7.7676 | -103.1991
7.7713 | -103.1889
7.7749 | -103.1858
7.7786 | -103.1941
7.7823 | -103.2151
7.7859 | -103.2457
7.7896 | -103.2804
7.7933 | -103.316
7.7969 | -103.3527
7.8006 | -103.3914
7.8043 | -103.4253
7.8079 | -103.4527
7.8116 | -103.4724
7.8152 | -103.4814
7.8189 | -103.4782
7.8226 | -103.4703
7.8262 | -103.4549
7.8299 | -103.4339
7.8336 | -103.405
7.8372 | -103.3775
7.8409 | -103.3504
7.8446 | -103.3212
7.8482 | -103.3015
7.8519 | -103.2823
7.8556 | -103.2643
7.8592 | -103.2521
7.8629 | -103.2501
7.8666 | -103.2527
7.8702 | -103.2609
7.8739 | -103.2743
7.8776 | -103.291
7.8812 | -103.3139
7.8849 | -103.3357
7.8886 | -103.3609
7.8922 | -103.385
7.8959 | -103.4025
7.8996 | -103.4182
7.9032 | -103.437
7.9069 | -103.4578
7.9106 | -103.4715
7.9142 | -103.4771
7.9179 | -103.4831
7.9216 | -103.4911
7.9252 | -103.4978
7.9289 | -103.5051
7.9326 | -103.5178
7.9362 | -103.5316
7.9399 | -103.5465
7.9435 | -103.5637
7.9472 | -103.5794
7.9509 | -103.5914
7.9545 | -103.5985
7.9582 | -103.6092
7.9619 | -103.62
7.9655 | -103.6236
7.9692 | -103.6227
7.9729 | -103.6207
7.9765 | -103.6191
7.9802 | -103.6164
7.9839 | -103.6175
7.9875 | -103.6162
7.9912 | -103.6157
7.9949 | -103.6215
7.9985 | -103.6293
8.0022 | -103.6391
8.0059 | -103.6482
8.0095 | -103.6615
8.0132 | -103.6777
8.0169 | -103.6963
8.0205 | -103.7113
8.0242 | -103.7254
8.0279 | -103.7416
8.0315 | -103.7623
8.0352 | -103.7777
8.0389 | -103.7937
8.0425 | -103.8107
8.0462 | -103.8255
8.0499 | -103.8357
8.0535 | -103.8412
8.0572 | -103.8413
8.0609 | -103.8369
8.0645 | -103.8314
8.0682 | -103.8244
8.0718 | -103.817
8.0755 | -103.7997
8.0792 | -103.7744
8.0828 | -103.7431
8.0865 | -103.7096
8.0902 | -103.6729
8.0938 | -103.6338
8.0975 | -103.5972
8.1012 | -103.561
8.1048 | -103.5262
8.1085 | -103.4993
8.1122 | -103.4795
8.1158 | -103.4631
8.1195 | -103.4512
8.1232 | -103.4477
8.1268 | -103.4494
8.1305 | -103.4575
8.1342 | -103.472
8.1378 | -103.4911
8.1415 | -103.5152
8.1452 | -103.5429
8.1488 | -103.5713
8.1525 | -103.5969
8.1562 | -103.6256
8.1598 | -103.6533
8.1635 | -103.6832
8.1672 | -103.7131
8.1708 | -103.7435
8.1745 | -103.7711
8.1782 | -103.7948
8.1818 | -103.8178
8.1855 | -103.8337
8.1891 | -103.8472
8.1928 | -103.8587
8.1965 | -103.8656
8.2001 | -103.872
8.2038 | -103.8804
8.2075 | -103.8815
8.2111 | -103.8787
8.2148 | -103.8735
8.2185 | -103.8632
8.2221 | -103.8438
8.2258 | -103.8212
8.2295 | -103.7951
8.2331 | -103.7672
8.2368 | -103.7377
8.2405 | -103.7061
8.2441 | -103.6812
8.2478 | -103.6604
8.2515 | -103.6403
8.2551 | -103.6265
8.2588 | -103.6224
8.2625 | -103.6261
8.2661 | -103.6349
8.2698 | -103.6494
8.2735 | -103.6693
8.2771 | -103.6909
8.2808 | -103.7151
8.2845 | -103.7382
8.2881 | -103.7652
8.2918 | -103.7937
8.2955 | -103.8216
8.2991 | -103.8489
8.3028 | -103.8725
8.3065 | -103.8903
8.3101 | -103.9029
8.3138 | -103.9133
8.3174 | -103.9226
8.3211 | -103.9239
8.3248 | -103.9232
8.3284 | -103.9219
8.3321 | -103.9147
8.3358 | -103.9036
8.3394 | -103.8877
8.3431 | -103.8729
8.3468 | -103.8538
8.3504 | -103.8339
8.3541 | -103.8106
8.3578 | -103.7906
8.3614 | -103.7726
8.3651 | -103.7513
8.3688 | -103.7313
8.3724 | -103.7152
8.3761 | -103.7029
8.3798 | -103.6884
8.3834 | -103.6789
8.3871 | -103.6742
8.3908 | -103.6721
8.3944 | -103.6711
8.3981 | -103.6722
8.4018 | -103.676
8.4054 | -103.6795
8.4091 | -103.6788
8.4128 | -103.6782
8.4164 | -103.6798
8.4201 | -103.6816
8.4238 | -103.6886
8.4274 | -103.7005
8.4311 | -103.7114
8.4348 | -103.7256
8.4384 | -103.744
8.4421 | -103.7683
8.4457 | -103.7974
8.4494 | -103.834
8.4531 | -103.8748
8.4567 | -103.9183
8.4604 | -103.9655
8.4641 | -104.0081
8.4677 | -104.0489
8.4714 | -104.089
8.4751 | -104.1237
8.4787 | -104.1548
8.4824 | -104.1793
8.4861 | -104.2001
8.4897 | -104.2153
8.4934 | -104.2228
8.4971 | -104.2263
8.5007 | -104.224
8.5044 | -104.2172
8.5081 | -104.1994
8.5117 | -104.1699
8.5154 | -104.1354
8.5191 | -104.0984
8.5227 | -104.0555
8.5264 | -104.007
8.5301 | -103.9599
8.5337 | -103.9153
8.5374 | -103.8631
8.5411 | -103.8059
8.5447 | -103.7543
8.5484 | -103.7089
8.5521 | -103.6698
8.5557 | -103.6363
8.5594 | -103.6074
8.5631 | -103.586
8.5667 | -103.5684
8.5704 | -103.5595
8.574 | -103.5564
8.5777 | -103.5614
8.5814 | -103.5762
8.585 | -103.5987
8.5887 | -103.6272
8.5924 | -103.6602
8.596 | -103.702
8.5997 | -103.7435
8.6034 | -103.7881
8.607 | -103.835
8.6107 | -103.8828
8.6144 | -103.9295
8.618 | -103.9782
8.6217 | -104.0224
8.6254 | -104.0603
8.629 | -104.0934
8.6327 | -104.12
8.6364 | -104.143
8.64 | -104.1608
8.6437 | -104.168
8.6474 | -104.1609
8.651 | -104.1457
8.6547 | -104.1204
8.6584 | -104.0872
8.662 | -104.0496
8.6657 | -104.0164
8.6694 | -103.9804
8.673 | -103.9358
8.6767 | -103.8893
8.6804 | -103.8436
8.684 | -103.8033
8.6877 | -103.7686
8.6913 | -103.7356
8.695 | -103.7094
8.6987 | -103.685
8.7023 | -103.6682
8.706 | -103.6558
8.7097 | -103.6445
8.7133 | -103.6348
8.717 | -103.6272
8.7207 | -103.6219
8.7243 | -103.6159
8.728 | -103.6148
8.7317 | -103.6199
8.7353 | -103.6289
8.739 | -103.6401
8.7427 | -103.6546
8.7463 | -103.6739
8.75 | -103.7001
8.7537 | -103.733
8.7573 | -103.7731
8.761 | -103.8194
8.7647 | -103.8704
8.7683 | -103.9225
8.772 | -103.9733
8.7757 | -104.0291
8.7793 | -104.0821
8.783 | -104.1347
8.7867 | -104.179
8.7903 | -104.2126
8.794 | -104.243
8.7977 | -104.2736
8.8013 | -104.3008
8.805 | -104.3213
8.8087 | -104.3406
8.8123 | -104.3535
8.816 | -104.3583
8.8196 | -104.3628
8.8233 | -104.3671
8.827 | -104.3698
8.8306 | -104.3671
8.8343 | -104.3643
8.838 | -104.3591
8.8416 | -104.3468
8.8453 | -104.3317
8.849 | -104.3191
8.8526 | -104.3028
8.8563 | -104.2851
8.86 | -104.2625
8.8636 | -104.2381
8.8673 | -104.2124
8.871 | -104.1886
8.8746 | -104.1652
8.8783 | -104.1454
8.882 | -104.1261
8.8856 | -104.1055
8.8893 | -104.0873
8.893 | -104.067
8.8966 | -104.0524
8.9003 | -104.041
8.904 | -104.0301
8.9076 | -104.0242
8.9113 | -104.0177
8.915 | -104.0151
8.9186 | -104.017
8.9223 | -104.0163
8.926 | -104.0131
8.9296 | -104.0154
8.9333 | -104.0242
8.9369 | -104.0271
8.9406 | -104.0334
8.9443 | -104.0435
8.9479 | -104.0532
8.9516 | -104.0589
8.9553 | -104.065
8.9589 | -104.075
8.9626 | -104.081
8.9663 | -104.0835
8.9699 | -104.0884
8.9736 | -104.0903
8.9773 | -104.0833
8.9809 | -104.0786
8.9846 | -104.073
8.9883 | -104.063
8.9919 | -104.0582
8.9956 | -104.0558
8.9993 | -104.0591
9.0029 | -104.0651
9.0066 | -104.0717
9.0103 | -104.0784
9.0139 | -104.0878
9.0176 | -104.0998
9.0213 | -104.1107
9.0249 | -104.1297
9.0286 | -104.1464
9.0323 | -104.1565
9.0359 | -104.1685
9.0396 | -104.1808
9.0433 | -104.1942
9.0469 | -104.2043
9.0506 | -104.2117
9.0543 | -104.2185
9.0579 | -104.2271
9.0616 | -104.2288
9.0652 | -104.2259
9.0689 | -104.2261
9.0726 | -104.2259
9.0762 | -104.2213
9.0799 | -104.2151
9.0836 | -104.2114
9.0872 | -104.2093
9.0909 | -104.2037
9.0946 | -104.1995
9.0982 | -104.1994
9.1019 | -104.1998
9.1056 | -104.1994
9.1092 | -104.199
9.1129 | -104.1987
9.1166 | -104.1956
9.1202 | -104.1979
9.1239 | -104.2004
9.1276 | -104.1983
9.1312 | -104.187
9.1349 | -104.1739
9.1386 | -104.1634
9.1422 | -104.1526
9.1459 | -104.1363
9.1496 | -104.1202
9.1532 | -104.1031
9.1569 | -104.0851
9.1606 | -104.0667
9.1642 | -104.0493
9.1679 | -104.0321
9.1716 | -104.0054
9.1752 | -103.981
9.1789 | -103.9567
9.1826 | -103.9372
9.1862 | -103.921
9.1899 | -103.9052
9.1935 | -103.8916
9.1972 | -103.8833
9.2009 | -103.881
9.2045 | -103.8832
9.2082 | -103.8866
9.2119 | -103.8957
9.2155 | -103.9098
9.2192 | -103.9244
9.2229 | -103.9381
9.2265 | -103.9573
9.2302 | -103.9799
9.2339 | -104.0014
9.2375 | -104.0199
9.2412 | -104.0415
9.2449 | -104.0648
9.2485 | -104.0838
9.2522 | -104.1028
9.2559 | -104.1159
9.2595 | -104.1262
9.2632 | -104.1371
9.2669 | -104.15
9.2705 | -104.1568
9.2742 | -104.1615
9.2779 | -104.1675
9.2815 | -104.1782
9.2852 | -104.1826
9.2889 | -104.185
9.2925 | -104.1864
9.2962 | -104.1849
9.2999 | -104.1866
9.3035 | -104.1931
9.3072 | -104.2014
9.3109 | -104.2077
9.3145 | -104.2184
9.3182 | -104.231
9.3218 | -104.2439
9.3255 | -104.2632
9.3292 | -104.2867
9.3328 | -104.3135
9.3365 | -104.3412
9.3402 | -104.3733
9.3438 | -104.4098
9.3475 | -104.4495
9.3512 | -104.4915
9.3548 | -104.5354
9.3585 | -104.5817
9.3622 | -104.6266
9.3658 | -104.6683
9.3695 | -104.7111
9.3732 | -104.7524
9.3768 | -104.7888
9.3805 | -104.8153
9.3842 | -104.8369
9.3878 | -104.8564
9.3915 | -104.8735
9.3952 | -104.8861
9.3988 | -104.896
9.4025 | -104.9015
9.4062 | -104.8996
9.4098 | -104.888
9.4135 | -104.8747
9.4172 | -104.8608
9.4208 | -104.8457
9.4245 | -104.8282
9.4282 | -104.8113
9.4318 | -104.793
9.4355 | -104.7691
9.4391 | -104.7447
9.4428 | -104.7263
9.4465 | -104.7049
9.4501 | -104.6798
9.4538 | -104.6582
9.4575 | -104.6383
9.4611 | -104.6189
9.4648 | -104.6006
9.4685 | -104.5821
9.4721 | -104.5693
9.4758 | -104.5586
9.4795 | -104.5506
9.4831 | -104.5445
9.4868 | -104.5391
9.4905 | -104.5366
9.4941 | -104.5354
9.4978 | -104.5352
9.5015 | -104.5313
9.5051 | -104.5279
9.5088 | -104.5283
9.5125 | -104.5275
9.5161 | -104.5249
9.5198 | -104.5304
9.5235 | -104.536
9.5271 | -104.5374
9.5308 | -104.5362
9.5345 | -104.5356
9.5381 | -104.5311
9.5418 | -104.5283
9.5455 | -104.5277
9.5491 | -104.5293
9.5528 | -104.5308
9.5565 | -104.5278
9.5601 | -104.524
9.5638 | -104.522
9.5674 | -104.5153
9.5711 | -104.507
9.5748 | -104.5032
9.5784 | -104.4933
9.5821 | -104.4777
9.5858 | -104.464
9.5894 | -104.4519
9.5931 | -104.4379
9.5968 | -104.426
9.6004 | -104.4083
9.6041 | -104.3949
9.6078 | -104.387
9.6114 | -104.3771
9.6151 | -104.3676
9.6188 | -104.3601
9.6224 | -104.354
9.6261 | -104.3549
9.6298 | -104.3594
9.6334 | -104.3655
9.6371 | -104.3748
9.6408 | -104.3896
9.6444 | -104.4051
9.6481 | -104.4208
9.6518 | -104.4322
9.6554 | -104.4445
9.6591 | -104.4552
9.6628 | -104.4608
9.6664 | -104.4702
9.6701 | -104.4808
9.6738 | -104.4925
9.6774 | -104.5002
9.6811 | -104.5092
9.6848 | -104.5157
9.6884 | -104.5159
9.6921 | -104.5096
9.6957 | -104.4998
9.6994 | -104.4896
9.7031 | -104.4763
9.7067 | -104.4592
9.7104 | -104.4437
9.7141 | -104.4278
9.7177 | -104.406
9.7214 | -104.3911
9.7251 | -104.377
9.7287 | -104.3632
9.7324 | -104.3538
9.7361 | -104.3478
9.7397 | -104.3446
9.7434 | -104.3517
9.7471 | -104.3566
9.7507 | -104.3686
9.7544 | -104.3865
9.7581 | -104.4068
9.7617 | -104.4252
9.7654 | -104.4431
9.7691 | -104.4603
9.7727 | -104.4799
9.7764 | -104.4989
9.7801 | -104.5146
9.7837 | -104.5323
9.7874 | -104.5478
9.7911 | -104.5624
9.7947 | -104.5809
9.7984 | -104.597
9.8021 | -104.6071
9.8057 | -104.6191
9.8094 | -104.6302
9.8131 | -104.6446
9.8167 | -104.6597
9.8204 | -104.6796
9.824 | -104.7016
9.8277 | -104.7237
9.8314 | -104.7516
9.835 | -104.781
9.8387 | -104.8128
9.8424 | -104.847
9.846 | -104.884
9.8497 | -104.9161
9.8534 | -104.944
9.857 | -104.9734
9.8607 | -105.0025
9.8644 | -105.0243
9.868 | -105.0429
9.8717 | -105.0562
9.8754 | -105.0626
9.879 | -105.0642
9.8827 | -105.0629
9.8864 | -105.0509
9.89 | -105.0402
9.8937 | -105.0257
9.8974 | -105.0066
9.901 | -104.9901
9.9047 | -104.9798
9.9084 | -104.9706
9.912 | -104.9601
9.9157 | -104.9492
9.9194 | -104.9415
9.923 | -104.937
9.9267 | -104.9325
9.9304 | -104.9313
9.934 | -104.9334
9.9377 | -104.9401
9.9413 | -104.9523
9.945 | -104.9682
9.9487 | -104.9838
9.9523 | -104.999
9.956 | -105.0102
9.9597 | -105.0195
9.9633 | -105.0241
9.967 | -105.0324
9.9707 | -105.0408
9.9743 | -105.0493
9.978 | -105.0577
9.9817 | -105.0561
9.9853 | -105.0475
9.989 | -105.0393
9.9927 | -105.0316
9.9963 | -105.0171
10 | -104.9977
10.0037 | -104.9773
10.0073 | -104.9512
10.011 | -104.9225
10.0147 | -104.8942
10.0183 | -104.8628
10.022 | -104.8279
10.0257 | -104.7896
10.0293 | -104.7541
10.033 | -104.7161
10.0367 | -104.6756
10.0403 | -104.6361
10.044 | -104.5981
10.0477 | -104.5686
10.0513 | -104.5353
10.055 | -104.5068
10.0587 | -104.4783
10.0623 | -104.4536
10.066 | -104.4315
10.0696 | -104.4098
10.0733 | -104.3941
10.077 | -104.3811
10.0806 | -104.3701
10.0843 | -104.3611
10.088 | -104.3519
10.0916 | -104.3459
10.0953 | -104.3455
10.099 | -104.3456
10.1026 | -104.3513
10.1063 | -104.3589
10.11 | -104.368
10.1136 | -104.3746
10.1173 | -104.3851
10.121 | -104.3924
10.1246 | -104.397
10.1283 | -104.4097
10.132 | -104.423
10.1356 | -104.4317
10.1393 | -104.44
10.143 | -104.447
10.1466 | -104.4518
10.1503 | -104.4558
10.154 | -104.4596
10.1576 | -104.4626
10.1613 | -104.4658
10.165 | -104.4734
10.1686 | -104.4832
10.1723 | -104.4945
10.176 | -104.5075
10.1796 | -104.5231
10.1833 | -104.5429
10.1869 | -104.568
10.1906 | -104.5965
10.1943 | -104.6256
10.1979 | -104.6532
10.2016 | -104.6813
10.2053 | -104.7134
10.2089 | -104.75
10.2126 | -104.7839
10.2163 | -104.8184
10.2199 | -104.8542
10.2236 | -104.8817
10.2273 | -104.9106
10.2309 | -104.9383
10.2346 | -104.9671
10.2383 | -104.9939
10.2419 | -105.0171
10.2456 | -105.031
10.2493 | -105.0474
10.2529 | -105.0567
10.2566 | -105.064
10.2603 | -105.0699
10.2639 | -105.0728
10.2676 | -105.0689
10.2713 | -105.0596
10.2749 | -105.0516
10.2786 | -105.0363
10.2823 | -105.0209
10.2859 | -105.0011
10.2896 | -104.9762
10.2933 | -104.9539
10.2969 | -104.9297
10.3006 | -104.9025
10.3043 | -104.8763
10.3079 | -104.8474
10.3116 | -104.8149
10.3152 | -104.7802
10.3189 | -104.7469
10.3226 | -104.7139
10.3262 | -104.6781
10.3299 | -104.6414
10.3336 | -104.6068
10.3372 | -104.5748
10.3409 | -104.5438
10.3446 | -104.516
10.3482 | -104.4957
10.3519 | -104.4789
10.3556 | -104.4637
10.3592 | -104.45
10.3629 | -104.4433
10.3666 | -104.4413
10.3702 | -104.4425
10.3739 | -104.4518
10.3776 | -104.4652
10.3812 | -104.4818
10.3849 | -104.4993
10.3886 | -104.5213
10.3922 | -104.5482
10.3959 | -104.5824
10.3996 | -104.6154
10.4032 | -104.6517
10.4069 | -104.6874
10.4106 | -104.7215
10.4142 | -104.7547
10.4179 | -104.7898
10.4216 | -104.825
10.4252 | -104.8541
10.4289 | -104.8801
10.4326 | -104.9005
10.4362 | -104.92
10.4399 | -104.9409
10.4435 | -104.957
10.4472 | -104.9676
10.4509 | -104.974
10.4545 | -104.972
10.4582 | -104.967
10.4619 | -104.9573
10.4655 | -104.9415
10.4692 | -104.9249
10.4729 | -104.9064
10.4765 | -104.8872
10.4802 | -104.8642
10.4839 | -104.8377
10.4875 | -104.8092
10.4912 | -104.7815
10.4949 | -104.7591
10.4985 | -104.7392
10.5022 | -104.7244
10.5059 | -104.7139
10.5095 | -104.7111
10.5132 | -104.7135
10.5169 | -104.7215
10.5205 | -104.7348
10.5242 | -104.7544
10.5279 | -104.7793
10.5315 | -104.8127
10.5352 | -104.8471
10.5389 | -104.8849
10.5425 | -104.9252
10.5462 | -104.9646
10.5499 | -105.0054
10.5535 | -105.0474
10.5572 | -105.0869
10.5609 | -105.1223
10.5645 | -105.1536
10.5682 | -105.1798
10.5718 | -105.2048
10.5755 | -105.203
10.5792 | -105.2176
10.5828 | -105.2276
10.5865 | -105.2321
10.5902 | -105.2322
10.5938 | -105.2242
10.5975 | -105.2104
10.6012 | -105.1961
10.6048 | -105.1816
10.6085 | -105.1684
10.6122 | -105.1484
10.6158 | -105.1286
10.6195 | -105.1129
10.6232 | -105.0932
10.6268 | -105.074
10.6305 | -105.0559
10.6342 | -105.0412
10.6378 | -105.0283
10.6415 | -105.0183
10.6452 | -105.0126
10.6488 | -105.007
10.6525 | -105.0011
10.6562 | -104.9964
10.6598 | -104.9937
10.6635 | -104.9905
10.6672 | -104.9885
10.6708 | -104.9857
10.6745 | -104.9842
10.6782 | -104.9795
10.6818 | -104.9743
10.6855 | -104.9667
10.6891 | -104.9609
10.6928 | -104.9531
10.6965 | -104.9421
10.7001 | -104.9303
10.7038 | -104.9221
10.7075 | -104.9105
10.7111 | -104.8991
10.7148 | -104.8863
10.7185 | -104.8727
10.7221 | -104.8614
10.7258 | -104.852
10.7295 | -104.8448
10.7331 | -104.8408
10.7368 | -104.8366
10.7405 | -104.8324
10.7441 | -104.8329
10.7478 | -104.8333
10.7515 | -104.8376
10.7551 | -104.8442
10.7588 | -104.8476
10.7625 | -104.8551
10.7661 | -104.8615
10.7698 | -104.8718
10.7735 | -104.8843
10.7771 | -104.8937
10.7808 | -104.9026
10.7845 | -104.9125
10.7881 | -104.922
10.7918 | -104.9291
10.7955 | -104.9364
10.7991 | -104.9428
10.8028 | -104.9469
10.8065 | -104.952
10.8101 | -104.9569
10.8138 | -104.9621
10.8174 | -104.9661
10.8211 | -104.9707
10.8248 | -104.975
10.8284 | -104.9801
10.8321 | -104.9843
10.8358 | -104.9878
10.8394 | -104.9916
10.8431 | -104.9983
10.8468 | -105.0027
10.8504 | -105.007
10.8541 | -105.0087
10.8578 | -105.0138
10.8614 | -105.015
10.8651 | -105.0184
10.8688 | -105.0181
10.8724 | -105.0165
10.8761 | -105.0172
10.8798 | -105.0151
10.8834 | -105.0117
10.8871 | -105.007
10.8908 | -105.0036
10.8944 | -105.0045
10.8981 | -105.003
10.9018 | -104.9964
10.9054 | -104.9913
10.9091 | -104.9866
10.9128 | -104.9818
10.9164 | -104.977
10.9201 | -104.9706
10.9238 | -104.9688
10.9274 | -104.9689
10.9311 | -104.97
10.9348 | -104.9702
10.9384 | -104.9722
10.9421 | -104.9759
10.9457 | -104.9804
10.9494 | -104.984
10.9531 | -104.9855
10.9567 | -104.9864
10.9604 | -104.9858
10.9641 | -104.9822
10.9677 | -104.9796
10.9714 | -104.9735
10.9751 | -104.9662
10.9787 | -104.9566
10.9824 | -104.9454
10.9861 | -104.9327
10.9897 | -104.9204
10.9934 | -104.9085
10.9971 | -104.8948
11.0007 | -104.8796
11.0044 | -104.8696
11.0081 | -104.859
11.0117 | -104.8455
11.0154 | -104.8303
11.0191 | -104.8193
11.0227 | -104.8107
11.0264 | -104.8058
11.0301 | -104.804
11.0337 | -104.8007
11.0374 | -104.8011
11.0411 | -104.8019
11.0447 | -104.8053
11.0484 | -104.8099
11.0521 | -104.8141
11.0557 | -104.8208
11.0594 | -104.8298
11.0631 | -104.8399
11.0667 | -104.8467
11.0704 | -104.8522
11.074 | -104.8602
11.0777 | -104.8679
11.0814 | -104.8725
11.085 | -104.8753
11.0887 | -104.8746
11.0924 | -104.8785
11.096 | -104.885
11.0997 | -104.8912
11.1034 | -104.9099
11.107 | -104.9241
11.1107 | -104.9389
11.1144 | -104.9586
11.118 | -104.9793
11.1217 | -105
11.1254 | -105.027
11.129 | -105.0541
11.1327 | -105.0812
11.1364 | -105.1111
11.14 | -105.1449
11.1437 | -105.1798
11.1474 | -105.2162
11.151 | -105.2502
11.1547 | -105.2834
11.1584 | -105.3174
11.162 | -105.3558
11.1657 | -105.3883
11.1694 | -105.4197
11.173 | -105.4562
11.1767 | -105.4856
11.1804 | -105.5106
11.184 | -105.5326
11.1877 | -105.5499
11.1913 | -105.5624
11.195 | -105.5718
11.1987 | -105.5774
11.2023 | -105.5825
11.206 | -105.5853
11.2097 | -105.5863
11.2133 | -105.5843
11.217 | -105.5845
11.2207 | -105.5833
11.2243 | -105.5762
11.228 | -105.5667
11.2317 | -105.5597
11.2353 | -105.5532
11.239 | -105.5453
11.2427 | -105.5377
11.2463 | -105.5274
11.25 | -105.5183
11.2537 | -105.5073
11.2573 | -105.4909
11.261 | -105.4727
11.2647 | -105.4513
11.2683 | -105.4283
11.272 | -105.4034
11.2757 | -105.3792
11.2793 | -105.3514
11.283 | -105.3211
11.2867 | -105.2941
11.2903 | -105.2635
11.294 | -105.2324
11.2977 | -105.204
11.3013 | -105.172
11.305 | -105.1425
11.3087 | -105.1156
11.3123 | -105.0895
11.316 | -105.0655
11.3196 | -105.0463
11.3233 | -105.0293
11.327 | -105.0108
11.3306 | -105.0015
11.3343 | -104.9943
11.338 | -104.9858
11.3416 | -104.9812
11.3453 | -104.9777
11.349 | -104.9777
11.3526 | -104.9788
11.3563 | -104.9814
11.36 | -104.9862
11.3636 | -104.9882
11.3673 | -104.9894
11.371 | -104.9949
11.3746 | -105.0011
11.3783 | -105.0057
11.382 | -105.0074
11.3856 | -105.0069
11.3893 | -105.0047
11.393 | -105.0046
11.3966 | -105.0008
11.4003 | -104.9907
11.404 | -104.9796
11.4076 | -104.965
11.4113 | -104.9498
11.415 | -104.9327
11.4186 | -104.9156
11.4223 | -104.8981
11.426 | -104.8796
11.4296 | -104.8638
11.4333 | -104.8492
11.4369 | -104.8357
11.4406 | -104.8227
11.4443 | -104.8115
11.4479 | -104.802
11.4516 | -104.7925
11.4553 | -104.7872
11.4589 | -104.7835
11.4626 | -104.7818
11.4663 | -104.7865
11.4699 | -104.791
11.4736 | -104.7952
11.4773 | -104.802
11.4809 | -104.8088
11.4846 | -104.8152
11.4883 | -104.8218
11.4919 | -104.8317
11.4956 | -104.8375
11.4993 | -104.8454
11.5029 | -104.8573
11.5066 | -104.868
11.5103 | -104.8799
11.5139 | -104.8905
11.5176 | -104.9045
11.5213 | -104.9206
11.5249 | -104.9324
11.5286 | -104.9464
11.5323 | -104.9625
11.5359 | -104.975
11.5396 | -104.9915
11.5433 | -105.01
11.5469 | -105.0263
11.5506 | -105.0443
11.5543 | -105.0652
11.5579 | -105.0844
11.5616 | -105.1048
11.5652 | -105.1262
11.5689 | -105.1462
11.5726 | -105.1698
11.5762 | -105.1928
11.5799 | -105.2146
11.5836 | -105.2317
11.5872 | -105.2498
11.5909 | -105.2713
11.5946 | -105.2897
11.5982 | -105.3056
11.6019 | -105.3219
11.6056 | -105.3384
11.6092 | -105.3547
11.6129 | -105.372
11.6166 | -105.3873
11.6202 | -105.4005
11.6239 | -105.4178
11.6276 | -105.4345
11.6312 | -105.4427
11.6349 | -105.4503
11.6386 | -105.4607
11.6422 | -105.4709
11.6459 | -105.4821
11.6496 | -105.4963
11.6532 | -105.5072
11.6569 | -105.5178
11.6606 | -105.5243
11.6642 | -105.5295
11.6679 | -105.537
11.6716 | -105.5402
11.6752 | -105.5393
11.6789 | -105.5326
11.6826 | -105.5293
11.6862 | -105.5231
11.6899 | -105.5133
11.6935 | -105.5091
11.6972 | -105.5054
11.7009 | -105.4994
11.7045 | -105.491
11.7082 | -105.4789
11.7119 | -105.4703
11.7155 | -105.4665
11.7192 | -105.4573
11.7229 | -105.446
11.7265 | -105.4363
11.7302 | -105.4262
11.7339 | -105.4206
11.7375 | -105.4174
11.7412 | -105.413
11.7449 | -105.4073
11.7485 | -105.401
11.7522 | -105.399
11.7559 | -105.3961
11.7595 | -105.391
11.7632 | -105.3881
11.7669 | -105.3865
11.7705 | -105.3841
11.7742 | -105.3814
11.7779 | -105.3807
11.7815 | -105.3784
11.7852 | -105.3718
11.7889 | -105.3632
11.7925 | -105.3523
11.7962 | -105.3388
11.7999 | -105.3262
11.8035 | -105.3145
11.8072 | -105.2993
11.8109 | -105.2827
11.8145 | -105.2686
11.8182 | -105.2501
11.8218 | -105.2322
11.8255 | -105.216
11.8292 | -105.1994
11.8328 | -105.1822
11.8365 | -105.1675
11.8402 | -105.152
11.8438 | -105.1397
11.8475 | -105.1338
11.8512 | -105.1319
11.8548 | -105.1257
11.8585 | -105.1222
11.8622 | -105.1213
11.8658 | -105.1183
11.8695 | -105.1165
11.8732 | -105.1192
11.8768 | -105.1247
11.8805 | -105.1317
11.8842 | -105.1396
11.8878 | -105.1469
11.8915 | -105.158
11.8952 | -105.1674
11.8988 | -105.1743
11.9025 | -105.1788
11.9062 | -105.1816
11.9098 | -105.1828
11.9135 | -105.184
11.9172 | -105.1875
11.9208 | -105.1918
11.9245 | -105.1932
11.9282 | -105.1926
11.9318 | -105.1904
11.9355 | -105.1865
11.9391 | -105.1845
11.9428 | -105.1857
11.9465 | -105.1871
11.9501 | -105.185
11.9538 | -105.1819
11.9575 | -105.1802
11.9611 | -105.1797
11.9648 | -105.1789
11.9685 | -105.1768
11.9721 | -105.1732
11.9758 | -105.1705
11.9795 | -105.1657
11.9831 | -105.1565
11.9868 | -105.15
11.9905 | -105.1485
11.9941 | -105.1471
11.9978 | -105.1477
12.0015 | -105.1501
12.0051 | -105.1577
12.0088 | -105.1666
12.0125 | -105.1752
12.0161 | -105.1875
12.0198 | -105.201
12.0235 | -105.2168
12.0271 | -105.2382
12.0308 | -105.2593
12.0345 | -105.2824
12.0381 | -105.3106
12.0418 | -105.3386
12.0455 | -105.3642
12.0491 | -105.391
12.0528 | -105.4188
12.0565 | -105.4457
12.0601 | -105.4718
12.0638 | -105.4986
12.0674 | -105.5262
12.0711 | -105.5514
12.0748 | -105.5751
12.0784 | -105.5949
12.0821 | -105.6155
12.0858 | -105.6325
12.0894 | -105.6454
12.0931 | -105.6553
12.0968 | -105.6649
12.1004 | -105.6718
12.1041 | -105.6758
12.1078 | -105.6783
12.1114 | -105.6797
12.1151 | -105.6769
12.1188 | -105.6707
12.1224 | -105.6643
12.1261 | -105.659
12.1298 | -105.656
12.1334 | -105.6482
12.1371 | -105.638
12.1408 | -105.6273
12.1444 | -105.6134
12.1481 | -105.596
12.1518 | -105.5816
12.1554 | -105.5708
12.1591 | -105.558
12.1628 | -105.5429
12.1664 | -105.5276
12.1701 | -105.5107
12.1738 | -105.4981
12.1774 | -105.4889
12.1811 | -105.4812
12.1848 | -105.4743
12.1884 | -105.4676
12.1921 | -105.4585
12.1957 | -105.4532
12.1994 | -105.4494
12.2031 | -105.4473
12.2067 | -105.4482
12.2104 | -105.4477
12.2141 | -105.4463
12.2177 | -105.4472
12.2214 | -105.4507
12.2251 | -105.4547
12.2287 | -105.4595
12.2324 | -105.4669
12.2361 | -105.4753
12.2397 | -105.4808
12.2434 | -105.487
12.2471 | -105.4945
12.2507 | -105.5068
12.2544 | -105.5217
12.2581 | -105.5334
12.2617 | -105.5441
12.2654 | -105.5565
12.2691 | -105.5696
12.2727 | -105.5808
12.2764 | -105.5953
12.2801 | -105.6041
12.2837 | -105.6093
12.2874 | -105.6128
12.2911 | -105.6167
12.2947 | -105.6216
12.2984 | -105.6217
12.3021 | -105.6164
12.3057 | -105.6104
12.3094 | -105.6069
12.3131 | -105.5997
12.3167 | -105.5873
12.3204 | -105.5754
12.324 | -105.5593
12.3277 | -105.5433
12.3314 | -105.5244
12.335 | -105.5047
12.3387 | -105.4872
12.3424 | -105.4711
12.346 | -105.4536
12.3497 | -105.4354
12.3534 | -105.4176
12.357 | -105.4006
12.3607 | -105.385
12.3644 | -105.3728
12.368 | -105.364
12.3717 | -105.3598
12.3754 | -105.3524
12.379 | -105.3455
12.3827 | -105.3431
12.3864 | -105.3414
12.39 | -105.3404
12.3937 | -105.3417
12.3974 | -105.3437
12.401 | -105.3443
12.4047 | -105.3437
12.4084 | -105.3414
12.412 | -105.3356
12.4157 | -105.333
12.4194 | -105.3322
12.423 | -105.3314
12.4267 | -105.3294
12.4304 | -105.3217
12.434 | -105.3128
12.4377 | -105.3067
12.4413 | -105.3046
12.445 | -105.2994
12.4487 | -105.2935
12.4523 | -105.2869
12.456 | -105.2801
12.4597 | -105.2802
12.4633 | -105.279
12.467 | -105.2756
12.4707 | -105.2747
12.4743 | -105.2752
12.478 | -105.2796
12.4817 | -105.2885
12.4853 | -105.2939
12.489 | -105.2972
12.4927 | -105.3056
12.4963 | -105.3162
12.5 | -105.3268
12.5037 | -105.3399
12.5073 | -105.3469
12.511 | -105.3539
12.5147 | -105.3633
12.5183 | -105.3737
12.522 | -105.3805
12.5257 | -105.3854
12.5293 | -105.3927
12.533 | -105.3949
12.5367 | -105.3936
12.5403 | -105.3923
12.544 | -105.3904
12.5477 | -105.3856
12.5513 | -105.3745
12.555 | -105.3683
12.5587 | -105.3609
12.5623 | -105.3467
12.566 | -105.3341
12.5696 | -105.3252
12.5733 | -105.3152
12.577 | -105.3057
12.5806 | -105.2965
12.5843 | -105.2906
12.588 | -105.287
12.5916 | -105.2888
12.5953 | -105.2898
12.599 | -105.2923
12.6026 | -105.2999
12.6063 | -105.3075
12.61 | -105.3183
12.6136 | -105.3314
12.6173 | -105.3416
12.621 | -105.3521
12.6246 | -105.366
12.6283 | -105.3796
12.632 | -105.3915
12.6356 | -105.4036
12.6393 | -105.4133
12.643 | -105.4219
12.6466 | -105.431
12.6503 | -105.4322
12.654 | -105.4305
12.6576 | -105.4263
12.6613 | -105.417
12.665 | -105.4059
12.6686 | -105.3941
12.6723 | -105.3812
12.676 | -105.37
12.6796 | -105.355
12.6833 | -105.3386
12.6869 | -105.3241
12.6906 | -105.3123
12.6943 | -105.3006
12.6979 | -105.2861
12.7016 | -105.2742
12.7053 | -105.2652
12.7089 | -105.2575
12.7126 | -105.2519
12.7163 | -105.2511
12.7199 | -105.2542
12.7236 | -105.2574
12.7273 | -105.2614
12.7309 | -105.2714
12.7346 | -105.2862
12.7383 | -105.2984
12.7419 | -105.3169
12.7456 | -105.3361
12.7493 | -105.3543
12.7529 | -105.3724
12.7566 | -105.3935
12.7603 | -105.4158
12.7639 | -105.4399
12.7676 | -105.4666
12.7713 | -105.4932
12.7749 | -105.5179
12.7786 | -105.5397
12.7823 | -105.5617
12.7859 | -105.5782
12.7896 | -105.5991
12.7933 | -105.6198
12.7969 | -105.6357
12.8006 | -105.653
12.8043 | -105.6673
12.8079 | -105.6784
12.8116 | -105.6861
12.8152 | -105.6942
12.8189 | -105.6995
12.8226 | -105.7028
12.8262 | -105.708
12.8299 | -105.7116
12.8336 | -105.7192
12.8372 | -105.7269
12.8409 | -105.7336
12.8446 | -105.7385
12.8482 | -105.7436
12.8519 | -105.7513
12.8556 | -105.7564
12.8592 | -105.762
12.8629 | -105.7697
12.8666 | -105.7742
12.8702 | -105.7808
12.8739 | -105.786
12.8776 | -105.7904
12.8812 | -105.7969
12.8849 | -105.8072
12.8886 | -105.814
12.8922 | -105.8177
12.8959 | -105.8202
12.8996 | -105.8198
12.9032 | -105.8176
12.9069 | -105.8145
12.9106 | -105.8122
12.9142 | -105.8106
12.9179 | -105.8056
12.9216 | -105.7988
12.9252 | -105.7913
12.9289 | -105.7788
12.9326 | -105.7702
12.9362 | -105.7635
12.9399 | -105.7539
12.9435 | -105.743
12.9472 | -105.7332
12.9509 | -105.7209
12.9545 | -105.7068
12.9582 | -105.6966
12.9619 | -105.6897
12.9655 | -105.6812
12.9692 | -105.6692
12.9729 | -105.6621
12.9765 | -105.652
12.9802 | -105.6403
12.9839 | -105.6335
12.9875 | -105.6246
12.9912 | -105.6148
12.9949 | -105.6075
12.9985 | -105.602
13.0022 | -105.595
13.0059 | -105.5863
13.0095 | -105.5773
13.0132 | -105.5671
13.0169 | -105.5543
13.0205 | -105.5403
13.0242 | -105.527
13.0279 | -105.5134
13.0315 | -105.5032
13.0352 | -105.4957
13.0389 | -105.4836
13.0425 | -105.4704
13.0462 | -105.4584
13.0499 | -105.4489
13.0535 | -105.4452
13.0572 | -105.4397
13.0609 | -105.4345
13.0645 | -105.4335
13.0682 | -105.4322
13.0718 | -105.4315
13.0755 | -105.4326
13.0792 | -105.4378
13.0828 | -105.448
13.0865 | -105.459
13.0902 | -105.4712
13.0938 | -105.4878
13.0975 | -105.5028
13.1012 | -105.5187
13.1048 | -105.5364
13.1085 | -105.5558
13.1122 | -105.5763
13.1158 | -105.5988
13.1195 | -105.6206
13.1232 | -105.6425
13.1268 | -105.6655
13.1305 | -105.6852
13.1342 | -105.7047
13.1378 | -105.7236
13.1415 | -105.7409
13.1452 | -105.7546
13.1488 | -105.7677
13.1525 | -105.7868
13.1562 | -105.805
13.1598 | -105.8191
13.1635 | -105.8335
13.1672 | -105.8472
13.1708 | -105.8571
13.1745 | -105.8676
13.1782 | -105.8752
13.1818 | -105.881
13.1855 | -105.8856
13.1891 | -105.8904
13.1928 | -105.8928
13.1965 | -105.896
13.2001 | -105.9022
13.2038 | -105.911
13.2075 | -105.9206
13.2111 | -105.9266
13.2148 | -105.932
13.2185 | -105.9344
13.2221 | -105.9384
13.2258 | -105.9426
13.2295 | -105.9488
13.2331 | -105.9549
13.2368 | -105.959
13.2405 | -105.9614
13.2441 | -105.9605
13.2478 | -105.9589
13.2515 | -105.9557
13.2551 | -105.9478
13.2588 | -105.9357
13.2625 | -105.9222
13.2661 | -105.9071
13.2698 | -105.8875
13.2735 | -105.8682
13.2771 | -105.8476
13.2808 | -105.8254
13.2845 | -105.7999
13.2881 | -105.7768
13.2918 | -105.7536
13.2955 | -105.7287
13.2991 | -105.7033
13.3028 | -105.6797
13.3065 | -105.6548
13.3101 | -105.6344
13.3138 | -105.6125
13.3174 | -105.5917
13.3211 | -105.5772
13.3248 | -105.563
13.3284 | -105.549
13.3321 | -105.5373
13.3358 | -105.5268
13.3394 | -105.5195
13.3431 | -105.5105
13.3468 | -105.5045
13.3504 | -105.4998
13.3541 | -105.4971
13.3578 | -105.4958
13.3614 | -105.4975
13.3651 | -105.4992
13.3688 | -105.5024
13.3724 | -105.5066
13.3761 | -105.5129
13.3798 | -105.5193
13.3834 | -105.5296
13.3871 | -105.5403
13.3908 | -105.5485
13.3944 | -105.555
13.3981 | -105.5611
13.4018 | -105.5711
13.4054 | -105.5831
13.4091 | -105.5939
13.4128 | -105.606
13.4164 | -105.6204
13.4201 | -105.6314
13.4238 | -105.643
13.4274 | -105.6551
13.4311 | -105.6656
13.4348 | -105.6794
13.4384 | -105.6911
13.4421 | -105.7034
13.4457 | -105.7163
13.4494 | -105.7306
13.4531 | -105.7456
13.4567 | -105.7568
13.4604 | -105.7681
13.4641 | -105.7772
13.4677 | -105.7808
13.4714 | -105.7841
13.4751 | -105.7891
13.4787 | -105.795
13.4824 | -105.8005
13.4861 | -105.8072
13.4897 | -105.8162
13.4934 | -105.8269
13.4971 | -105.8357
13.5007 | -105.8464
13.5044 | -105.8549
13.5081 | -105.8669
13.5117 | -105.8816
13.5154 | -105.8975
13.5191 | -105.9158
13.5227 | -105.9396
13.5264 | -105.9615
13.5301 | -105.9851
13.5337 | -106.0073
13.5374 | -106.0326
13.5411 | -106.0587
13.5447 | -106.0829
13.5484 | -106.1065
13.5521 | -106.1295
13.5557 | -106.1533
13.5594 | -106.1786
13.5631 | -106.202
13.5667 | -106.2217
13.5704 | -106.2396
13.574 | -106.2584
13.5777 | -106.2728
13.5814 | -106.2854
13.585 | -106.2982
13.5887 | -106.3083
13.5924 | -106.3177
13.596 | -106.3236
13.5997 | -106.3253
13.6034 | -106.3215
13.607 | -106.317
13.6107 | -106.3065
13.6144 | -106.2935
13.618 | -106.2832
13.6217 | -106.2726
13.6254 | -106.2575
13.629 | -106.2387
13.6327 | -106.2201
13.6364 | -106.2015
13.64 | -106.1804
13.6437 | -106.1568
13.6474 | -106.1337
13.651 | -106.1132
13.6547 | -106.0945
13.6584 | -106.0757
13.662 | -106.0599
13.6657 | -106.0411
13.6694 | -106.0235
13.673 | -106.009
13.6767 | -105.9931
13.6804 | -105.978
13.684 | -105.9657
13.6877 | -105.9533
13.6913 | -105.9431
13.695 | -105.9335
13.6987 | -105.9236
13.7023 | -105.9158
13.706 | -105.9088
13.7097 | -105.8975
13.7133 | -105.887
13.717 | -105.8764
13.7207 | -105.8678
13.7243 | -105.8609
13.728 | -105.855
13.7317 | -105.8489
13.7353 | -105.841
13.739 | -105.8334
13.7427 | -105.8265
13.7463 | -105.8227
13.75 | -105.8169
13.7537 | -105.8105
13.7573 | -105.8059
13.761 | -105.804
13.7647 | -105.8036
13.7683 | -105.8042
13.772 | -105.8028
13.7757 | -105.8034
13.7793 | -105.8038
13.783 | -105.8057
13.7867 | -105.8094
13.7903 | -105.812
13.794 | -105.8154
13.7977 | -105.8199
13.8013 | -105.8272
13.805 | -105.8324
13.8087 | -105.8382
13.8123 | -105.8453
13.816 | -105.851
13.8196 | -105.8562
13.8233 | -105.8648
13.827 | -105.8703
13.8306 | -105.8756
13.8343 | -105.8816
13.838 | -105.888
13.8416 | -105.8954
13.8453 | -105.9023
13.849 | -105.9076
13.8526 | -105.9112
13.8563 | -105.917
13.86 | -105.92
13.8636 | -105.924
13.8673 | -105.9265
13.871 | -105.9275
13.8746 | -105.9285
13.8783 | -105.9276
13.882 | -105.9267
13.8856 | -105.9256
13.8893 | -105.9212
13.893 | -105.9169
13.8966 | -105.9127
13.9003 | -105.908
13.904 | -105.9031
13.9076 | -105.8982
13.9113 | -105.8922
13.915 | -105.8832
13.9186 | -105.8739
13.9223 | -105.8631
13.926 | -105.8528
13.9296 | -105.8439
13.9333 | -105.8361
13.9369 | -105.8278
13.9406 | -105.8194
13.9443 | -105.8109
13.9479 | -105.8021
13.9516 | -105.7947
13.9553 | -105.7863
13.9589 | -105.7763
13.9626 | -105.7667
13.9663 | -105.7553
13.9699 | -105.7451
13.9736 | -105.7403
13.9773 | -105.7366
13.9809 | -105.7318
13.9846 | -105.7272
13.9883 | -105.7273
13.9919 | -105.7297
13.9956 | -105.7318
13.9993 | -105.735
14.0029 | -105.7383
14.0066 | -105.7419
14.0103 | -105.7465
14.0139 | -105.7549
14.0176 | -105.7639
14.0213 | -105.776
14.0249 | -105.7882
14.0286 | -105.8028
14.0323 | -105.8191
14.0359 | -105.8327
14.0396 | -105.8483
14.0433 | -105.8623
14.0469 | -105.8751
14.0506 | -105.8906
14.0543 | -105.9058
14.0579 | -105.9165
14.0616 | -105.9248
14.0652 | -105.9325
14.0689 | -105.9396
14.0726 | -105.9448
14.0762 | -105.9524
14.0799 | -105.9577
14.0836 | -105.9594
14.0872 | -105.9602
14.0909 | -105.9631
14.0946 | -105.9621
14.0982 | -105.9603
14.1019 | -105.9574
14.1056 | -105.953
14.1092 | -105.9499
14.1129 | -105.9469
14.1166 | -105.9464
14.1202 | -105.9445
14.1239 | -105.9408
14.1276 | -105.9373
14.1312 | -105.9347
14.1349 | -105.9345
14.1386 | -105.9363
14.1422 | -105.9379
14.1459 | -105.9379
14.1496 | -105.9415
14.1532 | -105.9503
14.1569 | -105.9589
14.1606 | -105.9699
14.1642 | -105.9807
14.1679 | -105.9942
14.1716 | -106.0103
14.1752 | -106.0229
14.1789 | -106.0362
14.1826 | -106.0541
14.1862 | -106.072
14.1899 | -106.0886
14.1935 | -106.1041
14.1972 | -106.1196
14.2009 | -106.136
14.2045 | -106.1513
14.2082 | -106.1653
14.2119 | -106.1774
14.2155 | -106.1828
14.2192 | -106.1862
14.2229 | -106.1862
14.2265 | -106.1846
14.2302 | -106.1818
14.2339 | -106.178
14.2375 | -106.1726
14.2412 | -106.1658
14.2449 | -106.1578
14.2485 | -106.1506
14.2522 | -106.1399
14.2559 | -106.128
14.2595 | -106.1197
14.2632 | -106.1113
14.2669 | -106.1026
14.2705 | -106.0939
14.2742 | -106.0861
14.2779 | -106.0852
14.2815 | -106.0761
14.2852 | -106.0678
14.2889 | -106.0618
14.2925 | -106.0578
14.2962 | -106.0547
14.2999 | -106.0528
14.3035 | -106.0505
14.3072 | -106.0492
14.3109 | -106.0486
14.3145 | -106.0476
14.3182 | -106.0494
14.3218 | -106.0519
14.3255 | -106.0509
14.3292 | -106.0483
14.3328 | -106.0466
14.3365 | -106.0424
14.3402 | -106.0399
14.3438 | -106.0331
14.3475 | -106.0278
14.3512 | -106.0252
14.3548 | -106.0207
14.3585 | -106.0117
14.3622 | -106.0035
14.3658 | -105.9943
14.3695 | -105.9863
14.3732 | -105.9757
14.3768 | -105.9618
14.3805 | -105.9483
14.3842 | -105.9331
14.3878 | -105.9155
14.3915 | -105.8989
14.3952 | -105.8818
14.3988 | -105.8677
14.4025 | -105.8542
14.4062 | -105.8354
14.4098 | -105.8201
14.4135 | -105.8073
14.4172 | -105.7994
14.4208 | -105.7942
14.4245 | -105.7906
14.4282 | -105.7882
14.4318 | -105.7858
14.4355 | -105.7875
14.4391 | -105.7918
14.4428 | -105.7955
14.4465 | -105.801
14.4501 | -105.8102
14.4538 | -105.8186
14.4575 | -105.8303
14.4611 | -105.8445
14.4648 | -105.8599
14.4685 | -105.8741
14.4721 | -105.8874
14.4758 | -105.8983
14.4795 | -105.9114
14.4831 | -105.9219
14.4868 | -105.9318
14.4905 | -105.9442
14.4941 | -105.955
14.4978 | -105.9629
14.5015 | -105.9685
14.5051 | -105.9717
14.5088 | -105.9781
14.5125 | -105.9822
14.5161 | -105.983
14.5198 | -105.9832
14.5235 | -105.9829
14.5271 | -105.9782
14.5308 | -105.9736
14.5345 | -105.9729
14.5381 | -105.9698
14.5418 | -105.963
14.5455 | -105.9557
14.5491 | -105.9452
14.5528 | -105.9376
14.5565 | -105.9259
14.5601 | -105.9112
14.5638 | -105.8975
14.5674 | -105.8852
14.5711 | -105.8756
14.5748 | -105.8672
14.5784 | -105.8591
14.5821 | -105.8533
14.5858 | -105.8476
14.5894 | -105.8404
14.5931 | -105.8364
14.5968 | -105.8333
14.6004 | -105.8301
14.6041 | -105.8285
14.6078 | -105.8249
14.6114 | -105.8222
14.6151 | -105.8248
14.6188 | -105.8267
14.6224 | -105.8285
14.6261 | -105.8283
14.6298 | -105.829
14.6334 | -105.8305
14.6371 | -105.8351
14.6408 | -105.8428
14.6444 | -105.8467
14.6481 | -105.8502
14.6518 | -105.8591
14.6554 | -105.8665
14.6591 | -105.8729
14.6628 | -105.8817
14.6664 | -105.8892
14.6701 | -105.8988
14.6738 | -105.9073
14.6774 | -105.9177
14.6811 | -105.9299
14.6848 | -105.9444
14.6884 | -105.9556
14.6921 | -105.9679
14.6957 | -105.9808
14.6994 | -105.9927
14.7031 | -106.0039
14.7067 | -106.0164
14.7104 | -106.0279
14.7141 | -106.0391
14.7177 | -106.0464
14.7214 | -106.0532
14.7251 | -106.062
14.7287 | -106.0697
14.7324 | -106.0767
14.7361 | -106.0818
14.7397 | -106.0807
14.7434 | -106.0777
14.7471 | -106.0767
14.7507 | -106.071
14.7544 | -106.0627
14.7581 | -106.0529
14.7617 | -106.0412
14.7654 | -106.0284
14.7691 | -106.0117
14.7727 | -105.9911
14.7764 | -105.9676
14.7801 | -105.9424
14.7837 | -105.9168
14.7874 | -105.8915
14.7911 | -105.8674
14.7947 | -105.847
14.7984 | -105.8288
14.8021 | -105.8087
14.8057 | -105.8007
14.8094 | -105.786
14.8131 | -105.7722
14.8167 | -105.7595
14.8204 | -105.7515
14.824 | -105.7451
14.8277 | -105.7362
14.8314 | -105.728
14.835 | -105.7223
14.8387 | -105.7184
14.8424 | -105.7162
14.846 | -105.7129
14.8497 | -105.7093
14.8534 | -105.7083
14.857 | -105.7093
14.8607 | -105.7107
14.8644 | -105.7122
14.868 | -105.7157
14.8717 | -105.718
14.8754 | -105.722
14.879 | -105.7265
14.8827 | -105.727
14.8864 | -105.7283
14.89 | -105.7315
14.8937 | -105.7307
14.8974 | -105.7301
14.901 | -105.7347
14.9047 | -105.7411
14.9084 | -105.7476
14.912 | -105.7532
14.9157 | -105.7593
14.9194 | -105.7637
14.923 | -105.7701
14.9267 | -105.7787
14.9304 | -105.7893
14.934 | -105.8022
14.9377 | -105.8131
14.9413 | -105.8266
14.945 | -105.8432
14.9487 | -105.8584
14.9523 | -105.876
14.956 | -105.8975
14.9597 | -105.9194
14.9633 | -105.9435
14.967 | -105.9705
14.9707 | -105.9941
14.9743 | -106.0175
14.978 | -106.0425
14.9817 | -106.0685
14.9853 | -106.0941
14.989 | -106.1216
14.9927 | -106.1459
14.9963 | -106.1663
15 | -106.1859
15.0037 | -106.2037
15.0073 | -106.2176
15.011 | -106.2272
15.0147 | -106.239
15.0183 | -106.249
15.022 | -106.2538
15.0257 | -106.2592
15.0293 | -106.2625
15.033 | -106.2655
15.0367 | -106.2637
15.0403 | -106.2583
15.044 | -106.2558
15.0477 | -106.2475
15.0513 | -106.2356
15.055 | -106.2252
15.0587 | -106.2129
15.0623 | -106.2027
15.066 | -106.1922
15.0696 | -106.1769
15.0733 | -106.164
15.077 | -106.1513
15.0806 | -106.1378
15.0843 | -106.1263
15.088 | -106.1161
15.0916 | -106.1039
15.0953 | -106.0927
15.099 | -106.0871
15.1026 | -106.0853
15.1063 | -106.083
15.11 | -106.0834
15.1136 | -106.0814
15.1173 | -106.0803
15.121 | -106.0812
15.1246 | -106.0799
15.1283 | -106.0811
15.132 | -106.0876
15.1356 | -106.0963
15.1393 | -106.1047
15.143 | -106.1123
15.1466 | -106.1187
15.1503 | -106.1251
15.154 | -106.1318
15.1576 | -106.1401
15.1613 | -106.146
15.165 | -106.1568
15.1686 | -106.1649
15.1723 | -106.1678
15.176 | -106.1733
15.1796 | -106.1803
15.1833 | -106.1835
15.1869 | -106.1896
15.1906 | -106.1952
15.1943 | -106.1999
15.1979 | -106.2072
15.2016 | -106.2118
15.2053 | -106.2117
15.2089 | -106.2145
15.2126 | -106.2192
15.2163 | -106.2225
15.2199 | -106.2259
15.2236 | -106.2292
15.2273 | -106.2358
15.2309 | -106.2386
15.2346 | -106.2401
15.2383 | -106.2439
15.2419 | -106.2474
15.2456 | -106.2498
15.2493 | -106.252
15.2529 | -106.2569
15.2566 | -106.2647
15.2603 | -106.2704
15.2639 | -106.2788
15.2676 | -106.2867
15.2713 | -106.2874
15.2749 | -106.2902
15.2786 | -106.2919
15.2823 | -106.294
15.2859 | -106.2976
15.2896 | -106.2972
15.2933 | -106.2989
15.2969 | -106.3008
15.3006 | -106.2995
15.3043 | -106.2969
15.3079 | -106.2943
15.3116 | -106.2897
15.3152 | -106.2821
15.3189 | -106.2724
15.3226 | -106.2612
15.3262 | -106.2509
15.3299 | -106.2394
15.3336 | -106.2184
15.3372 | -106.2043
15.3409 | -106.1887
15.3446 | -106.175
15.3482 | -106.1596
15.3519 | -106.1399
15.3556 | -106.1205
15.3592 | -106.1026
15.3629 | -106.0858
15.3666 | -106.069
15.3702 | -106.0524
15.3739 | -106.0369
15.3776 | -106.0207
15.3812 | -106.0071
15.3849 | -105.9941
15.3886 | -105.9825
15.3922 | -105.9695
15.3959 | -105.9572
15.3996 | -105.9455
15.4032 | -105.9381
15.4069 | -105.9331
15.4106 | -105.9293
15.4142 | -105.9268
15.4179 | -105.9259
15.4216 | -105.9264
15.4252 | -105.9291
15.4289 | -105.9303
15.4326 | -105.9331
15.4362 | -105.9367
15.4399 | -105.9394
15.4435 | -105.9444
15.4472 | -105.9484
15.4509 | -105.9518
15.4545 | -105.9551
15.4582 | -105.9581
15.4619 | -105.9611
15.4655 | -105.9636
15.4692 | -105.9648
15.4729 | -105.9655
15.4765 | -105.9642
15.4802 | -105.9606
15.4839 | -105.9576
15.4875 | -105.9562
15.4912 | -105.9565
15.4949 | -105.9528
15.4985 | -105.9492
15.5022 | -105.9447
15.5059 | -105.9383
15.5095 | -105.9334
15.5132 | -105.9301
15.5169 | -105.9276
15.5205 | -105.9253
15.5242 | -105.9246
15.5279 | -105.9237
15.5315 | -105.9241
15.5352 | -105.9266
15.5389 | -105.9292
15.5425 | -105.9328
15.5462 | -105.9382
15.5499 | -105.9477
15.5535 | -105.957
15.5572 | -105.9659
15.5609 | -105.9761
15.5645 | -105.9894
15.5682 | -106.0034
15.5718 | -106.0168
15.5755 | -106.0323
15.5792 | -106.0464
15.5828 | -106.0586
15.5865 | -106.0686
15.5902 | -106.0803
15.5938 | -106.0922
15.5975 | -106.1049
15.6012 | -106.1161
15.6048 | -106.1262
15.6085 | -106.1381
15.6122 | -106.1501
15.6158 | -106.1633
15.6195 | -106.1728
15.6232 | -106.1812
15.6268 | -106.1845
15.6305 | -106.1865
15.6342 | -106.1893
15.6378 | -106.1901
15.6415 | -106.1875
15.6452 | -106.1843
15.6488 | -106.1813
15.6525 | -106.1765
15.6562 | -106.1704
15.6598 | -106.1634
15.6635 | -106.1544
15.6672 | -106.1415
15.6708 | -106.1296
15.6745 | -106.1191
15.6782 | -106.1084
15.6818 | -106.0947
15.6855 | -106.0793
15.6891 | -106.0674
15.6928 | -106.0569
15.6965 | -106.0469
15.7001 | -106.0384
15.7038 | -106.0298
15.7075 | -106.0209
15.7111 | -106.0133
15.7148 | -106.0075
15.7185 | -106.0035
15.7221 | -106
15.7258 | -105.9973
15.7295 | -105.9941
15.7331 | -105.9912
15.7368 | -105.9872
15.7405 | -105.9871
15.7441 | -105.989
15.7478 | -105.9899
15.7515 | -105.9903
15.7551 | -105.9928
15.7588 | -105.9939
15.7625 | -105.9942
15.7661 | -105.9951
15.7698 | -105.9953
15.7735 | -105.9963
15.7771 | -105.9971
15.7808 | -105.9961
15.7845 | -105.9963
15.7881 | -105.9954
15.7918 | -105.9941
15.7955 | -105.9943
15.7991 | -105.9951
15.8028 | -105.9959
15.8065 | -105.9969
15.8101 | -105.9999
15.8138 | -106.0002
15.8174 | -106.0005
15.8211 | -105.9998
15.8248 | -106.0005
15.8284 | -105.9982
15.8321 | -105.9914
15.8358 | -105.9844
15.8394 | -105.978
15.8431 | -105.9735
15.8468 | -105.9687
15.8504 | -105.9661
15.8541 | -105.9637
15.8578 | -105.9575
15.8614 | -105.953
15.8651 | -105.953
15.8688 | -105.9479
15.8724 | -105.9458
15.8761 | -105.9421
15.8798 | -105.9398
15.8834 | -105.9358
15.8871 | -105.9297
15.8908 | -105.925
15.8944 | -105.9212
15.8981 | -105.9182
15.9018 | -105.9135
15.9054 | -105.91
15.9091 | -105.9092
15.9128 | -105.9071
15.9164 | -105.9043
15.9201 | -105.899
15.9238 | -105.894
15.9274 | -105.8901
15.9311 | -105.8853
15.9348 | -105.8794
15.9384 | -105.8723
15.9421 | -105.8658
15.9457 | -105.8613
15.9494 | -105.8571
15.9531 | -105.852
15.9567 | -105.8506
15.9604 | -105.8502
15.9641 | -105.849
15.9677 | -105.8468
15.9714 | -105.8444
15.9751 | -105.8456
15.9787 | -105.8473
15.9824 | -105.8479
15.9861 | -105.8486
15.9897 | -105.8516
15.9934 | -105.8574
15.9971 | -105.8639
16.0007 | -105.8706
16.0044 | -105.8741
16.0081 | -105.8821
16.0117 | -105.8895
16.0154 | -105.8969
16.0191 | -105.9053
16.0227 | -105.9146
16.0264 | -105.9234
16.0301 | -105.9335
16.0337 | -105.9433
16.0374 | -105.9519
16.0411 | -105.9619
16.0447 | -105.9716
16.0484 | -105.9806
16.0521 | -105.9863
16.0557 | -105.995
16.0594 | -106.0087
16.0631 | -106.0236
16.0667 | -106.0336
16.0704 | -106.0414
16.074 | -106.0496
16.0777 | -106.0577
16.0814 | -106.0627
16.085 | -106.0694
16.0887 | -106.0771
16.0924 | -106.0812
16.096 | -106.0845
16.0997 | -106.0898
16.1034 | -106.0935
16.107 | -106.0974
16.1107 | -106.0992
16.1144 | -106.1011
16.118 | -106.1039
16.1217 | -106.108
16.1254 | -106.1106
16.129 | -106.1125
16.1327 | -106.1151
16.1364 | -106.1124
16.14 | -106.1083
16.1437 | -106.1085
16.1474 | -106.109
16.151 | -106.1103
16.1547 | -106.1119
16.1584 | -106.1142
16.162 | -106.1185
16.1657 | -106.1245
16.1694 | -106.1315
16.173 | -106.1378
16.1767 | -106.1437
16.1804 | -106.1511
16.184 | -106.1596
16.1877 | -106.1678
16.1913 | -106.175
16.195 | -106.1834
16.1987 | -106.192
16.2023 | -106.2026
16.206 | -106.2111
16.2097 | -106.2201
16.2133 | -106.2296
16.217 | -106.2396
16.2207 | -106.249
16.2243 | -106.2571
16.228 | -106.2635
16.2317 | -106.2703
16.2353 | -106.2829
16.239 | -106.2928
16.2427 | -106.3013
16.2463 | -106.3091
16.25 | -106.3163
16.2537 | -106.3221
16.2573 | -106.3253
16.261 | -106.3285
16.2647 | -106.332
16.2683 | -106.3359
16.272 | -106.3391
16.2757 | -106.339
16.2793 | -106.3381
16.283 | -106.3391
16.2867 | -106.3398
16.2903 | -106.3378
16.294 | -106.3366
16.2977 | -106.3404
16.3013 | -106.3414
16.305 | -106.3405
16.3087 | -106.3416
16.3123 | -106.3464
16.316 | -106.3528
16.3196 | -106.3584
16.3233 | -106.3665
16.327 | -106.3752
16.3306 | -106.3838
16.3343 | -106.3925
16.338 | -106.402
16.3416 | -106.4129
16.3453 | -106.4247
16.349 | -106.4369
16.3526 | -106.4484
16.3563 | -106.4624
16.36 | -106.4757
16.3636 | -106.4876
16.3673 | -106.5033
16.371 | -106.5205
16.3746 | -106.5354
16.3783 | -106.5471
16.382 | -106.5589
16.3856 | -106.5708
16.3893 | -106.5797
16.393 | -106.5771
16.3966 | -106.5837
16.4003 | -106.5874
16.404 | -106.5897
16.4076 | -106.593
16.4113 | -106.5914
16.415 | -106.5863
16.4186 | -106.5835
16.4223 | -106.5789
16.426 | -106.5742
16.4296 | -106.5688
16.4333 | -106.5612
16.4369 | -106.5534
16.4406 | -106.5465
16.4443 | -106.5395
16.4479 | -106.5325
16.4516 | -106.5228
16.4553 | -106.5126
16.4589 | -106.5016
16.4626 | -106.4917
16.4663 | -106.4788
16.4699 | -106.4665
16.4736 | -106.4546
16.4773 | -106.4409
16.4809 | -106.4292
16.4846 | -106.4181
16.4883 | -106.4057
16.4919 | -106.3942
16.4956 | -106.3838
16.4993 | -106.3734
16.5029 | -106.3621
16.5066 | -106.3488
16.5103 | -106.337
16.5139 | -106.3279
16.5176 | -106.321
16.5213 | -106.3117
16.5249 | -106.3029
16.5286 | -106.2939
16.5323 | -106.2829
16.5359 | -106.2737
16.5396 | -106.2661
16.5433 | -106.2578
16.5469 | -106.2497
16.5506 | -106.242
16.5543 | -106.2347
16.5579 | -106.2276
16.5616 | -106.2197
16.5652 | -106.2123
16.5689 | -106.2059
16.5726 | -106.1994
16.5762 | -106.1935
16.5799 | -106.1895
16.5836 | -106.1876
16.5872 | -106.1865
16.5909 | -106.1833
16.5946 | -106.1779
16.5982 | -106.1761
16.6019 | -106.1779
16.6056 | -106.1788
16.6092 | -106.1798
16.6129 | -106.1828
16.6166 | -106.1875
16.6202 | -106.193
16.6239 | -106.1976
16.6276 | -106.2055
16.6312 | -106.2132
16.6349 | -106.2188
16.6386 | -106.2285
16.6422 | -106.2374
16.6459 | -106.2446
16.6496 | -106.2552
16.6532 | -106.2674
16.6569 | -106.2778
16.6606 | -106.2879
16.6642 | -106.2983
16.6679 | -106.309
16.6716 | -106.3192
16.6752 | -106.3287
16.6789 | -106.3384
16.6826 | -106.3485
16.6862 | -106.3604
16.6899 | -106.372
16.6935 | -106.3844
16.6972 | -106.3975
16.7009 | -106.4081
16.7045 | -106.4215
16.7082 | -106.4334
16.7119 | -106.4451
16.7155 | -106.4582
16.7192 | -106.4699
16.7229 | -106.4788
16.7265 | -106.4876
16.7302 | -106.4975
16.7339 | -106.5039
16.7375 | -106.5132
16.7412 | -106.5217
16.7449 | -106.5303
16.7485 | -106.5366
16.7522 | -106.5399
16.7559 | -106.5403
16.7595 | -106.5409
16.7632 | -106.537
16.7669 | -106.532
16.7705 | -106.5276
16.7742 | -106.5237
16.7779 | -106.5179
16.7815 | -106.5089
16.7852 | -106.4994
16.7889 | -106.4875
16.7925 | -106.4754
16.7962 | -106.4624
16.7999 | -106.4523
16.8035 | -106.4424
16.8072 | -106.4286
16.8108 | -106.4139
16.8145 | -106.4006
16.8182 | -106.3851
16.8218 | -106.3705
16.8255 | -106.3557
16.8292 | -106.3442
16.8328 | -106.3368
16.8365 | -106.3284
16.8402 | -106.3164
16.8438 | -106.3058
16.8475 | -106.306
16.8512 | -106.3062
16.8548 | -106.3064
16.8585 | -106.3066
16.8622 | -106.3068
16.8658 | -106.3069
16.8695 | -106.3071
16.8732 | -106.3073
16.8768 | -106.3075
16.8805 | -106.3077
16.8842 | -106.3079
16.8878 | -106.3081
16.8915 | -106.3082
16.8952 | -106.3084
16.8988 | -106.3086
16.9025 | -106.3088
16.9062 | -106.309
16.9098 | -106.3092
16.9135 | -106.3093
16.9172 | -106.3095
16.9208 | -106.3097
16.9245 | -106.3099
16.9282 | -106.3101
16.9318 | -106.3103
16.9355 | -106.3105
16.9392 | -106.3106
16.9428 | -106.3108
16.9465 | -106.311
16.9501 | -106.3112
16.9538 | -106.3114
16.9575 | -106.3116
16.9611 | -106.3117
16.9648 | -106.3119
16.9685 | -106.3121
16.9721 | -106.3123
16.9758 | -106.3125
16.9795 | -106.3127
16.9831 | -106.3129
16.9868 | -106.313
16.9905 | -106.3132
16.9941 | -106.3134
16.9978 | -106.3136
17.0015 | -106.3138
17.0051 | -106.314
17.0088 | -106.3141
17.0125 | -106.3143
17.0161 | -106.3145
17.0198 | -106.3147
17.0235 | -106.3149
17.0271 | -106.3151
17.0308 | -106.3153
17.0345 | -106.3154
17.0381 | -106.3156
17.0418 | -106.3158
17.0455 | -106.316
17.0491 | -106.3162
17.0528 | -106.3164
17.0565 | -106.3166
17.0601 | -106.3167
17.0638 | -106.3169
17.0674 | -106.3171
17.0711 | -106.3173
17.0748 | -106.3175
17.0784 | -106.3177
17.0821 | -106.3178
17.0858 | -106.318
17.0894 | -106.3182
17.0931 | -106.3184
17.0968 | -106.3186
17.1004 | -106.3188
17.1041 | -106.319
17.1078 | -106.3191
17.1114 | -106.3193
17.1151 | -106.3195
17.1188 | -106.3197
17.1224 | -106.3199
17.1261 | -106.3201
17.1298 | -106.3202
17.1334 | -106.3204
17.1371 | -106.3206
17.1408 | -106.3208
17.1444 | -106.321
17.1481 | -106.3212
17.1518 | -106.3214
17.1554 | -106.3215
17.1591 | -106.3217
17.1628 | -106.3219
17.1664 | -106.3221
17.1701 | -106.3223
17.1738 | -106.3225
17.1774 | -106.3226
17.1811 | -106.3228
17.1848 | -106.323
17.1884 | -106.3232
17.1921 | -106.3234
17.1957 | -106.3236
17.1994 | -106.3238
17.2031 | -106.3239
17.2067 | -106.3241
17.2104 | -106.3243
17.2141 | -106.3245
17.2177 | -106.3247
17.2214 | -106.3249
17.2251 | -106.3252
17.2287 | -106.3338
17.2324 | -106.3435
17.2361 | -106.3515
17.2397 | -106.3595
17.2434 | -106.3662
17.2471 | -106.3705
17.2507 | -106.3755
17.2544 | -106.3773
17.2581 | -106.3789
17.2617 | -106.3791
17.2654 | -106.38
17.2691 | -106.3801
17.2727 | -106.3753
17.2764 | -106.3696
17.2801 | -106.3625
17.2837 | -106.3566
17.2874 | -106.3511
17.2911 | -106.3438
17.2947 | -106.336
17.2984 | -106.3291
17.3021 | -106.3229
17.3057 | -106.3167
17.3094 | -106.3062
17.3131 | -106.2961
17.3167 | -106.2859
17.3204 | -106.2718
17.324 | -106.2571
17.3277 | -106.2448
17.3314 | -106.2323
17.335 | -106.2183
17.3387 | -106.2066
17.3424 | -106.1937
17.346 | -106.1793
17.3497 | -106.168
17.3534 | -106.1556
17.357 | -106.1404
17.3607 | -106.1274
17.3644 | -106.1176
17.368 | -106.1056
17.3717 | -106.0937
17.3754 | -106.082
17.379 | -106.0699
17.3827 | -106.0603
17.3864 | -106.0548
17.39 | -106.0485
17.3937 | -106.0412
17.3974 | -106.0351
17.401 | -106.0314
17.4047 | -106.0293
17.4084 | -106.0269
17.412 | -106.025
17.4157 | -106.0251
17.4194 | -106.0267
17.423 | -106.0279
17.4267 | -106.0336
17.4304 | -106.0386
17.434 | -106.0455
17.4377 | -106.0551
17.4413 | -106.0635
17.445 | -106.0723
17.4487 | -106.0894
17.4523 | -106.1006
17.456 | -106.1136
17.4597 | -106.1266
17.4633 | -106.1385
17.467 | -106.1521
17.4707 | -106.1661
17.4743 | -106.1801
17.478 | -106.1948
17.4817 | -106.2089
17.4853 | -106.223
17.489 | -106.2357
17.4927 | -106.2509
17.4963 | -106.2653
17.5 | -106.2763
17.5037 | -106.2862
17.5073 | -106.2952
17.511 | -106.3035
17.5147 | -106.3098
17.5183 | -106.3153
17.522 | -106.3235
17.5257 | -106.3285
17.5293 | -106.3285
17.533 | -106.3259
17.5367 | -106.3214
17.5403 | -106.3166
17.544 | -106.3105
17.5477 | -106.3019
17.5513 | -106.2926
17.555 | -106.2829
17.5587 | -106.2758
17.5623 | -106.2672
17.566 | -106.2551
17.5696 | -106.2426
17.5733 | -106.2302
17.577 | -106.218
17.5806 | -106.2049
17.5843 | -106.1931
17.588 | -106.1801
17.5916 | -106.1688
17.5953 | -106.1568
17.599 | -106.1426
17.6026 | -106.1312
17.6063 | -106.1209
17.61 | -106.1086
17.6136 | -106.0952
17.6173 | -106.0833
17.621 | -106.073
17.6246 | -106.0632
17.6283 | -106.0545
17.632 | -106.0477
17.6356 | -106.0403
17.6393 | -106.0335
17.643 | -106.0278
17.6466 | -106.0245
17.6503 | -106.0204
17.654 | -106.0171
17.6576 | -106.0155
17.6613 | -106.0145
17.665 | -106.0145
17.6686 | -106.0151
17.6723 | -106.0188
17.676 | -106.0227
17.6796 | -106.028
17.6833 | -106.0333
17.6869 | -106.0383
17.6906 | -106.047
17.6943 | -106.0567
17.6979 | -106.0665
17.7016 | -106.076
17.7053 | -106.086
17.7089 | -106.0963
17.7126 | -106.1062
17.7163 | -106.1187
17.7199 | -106.1317
17.7236 | -106.1451
17.7273 | -106.1564
17.7309 | -106.1647
17.7346 | -106.1727
17.7383 | -106.1832
17.7419 | -106.1955
17.7456 | -106.2024
17.7493 | -106.2098
17.7529 | -106.219
17.7566 | -106.2272
17.7603 | -106.2316
17.7639 | -106.2336
17.7676 | -106.2365
17.7713 | -106.2385
17.7749 | -106.2357
17.7786 | -106.2329
17.7823 | -106.2322
17.7859 | -106.2292
17.7896 | -106.2243
17.7933 | -106.2172
17.7969 | -106.2086
17.8006 | -106.1973
17.8043 | -106.1831
17.8079 | -106.1699
17.8116 | -106.1535
17.8152 | -106.1364
17.8189 | -106.1227
17.8226 | -106.109
17.8262 | -106.0922
17.8299 | -106.0737
17.8336 | -106.0592
17.8372 | -106.0454
17.8409 | -106.0323
17.8446 | -106.0189
17.8482 | -106.0069
17.8519 | -105.9969
17.8556 | -105.9857
17.8592 | -105.9746
17.8629 | -105.9638
17.8666 | -105.9577
17.8702 | -105.9521
17.8739 | -105.944
17.8776 | -105.9387
17.8812 | -105.9335
17.8849 | -105.9287
17.8886 | -105.9247
17.8922 | -105.9218
17.8959 | -105.9217
17.8996 | -105.9212
17.9032 | -105.9223
17.9069 | -105.9251
17.9106 | -105.9279
17.9142 | -105.9315
17.9179 | -105.9353
17.9216 | -105.9383
17.9252 | -105.9403
17.9289 | -105.9417
17.9326 | -105.9443
17.9362 | -105.9474
17.9399 | -105.9514
17.9435 | -105.9545
17.9472 | -105.9581
17.9509 | -105.9619
17.9545 | -105.9653
17.9582 | -105.9695
17.9619 | -105.9744
17.9655 | -105.9776
17.9692 | -105.9794
17.9729 | -105.9787
17.9765 | -105.9792
17.9802 | -105.9809
17.9839 | -105.9819
17.9875 | -105.9844
17.9912 | -105.9893
17.9949 | -105.9928
17.9985 | -105.9942
18.0022 | -105.9965
18.0059 | -105.9969
18.0095 | -105.9962
18.0132 | -105.9958
18.0169 | -105.9984
18.0205 | -106.0026
18.0242 | -106.0029
18.0279 | -106.0025
18.0315 | -106.0033
18.0352 | -106.0048
18.0389 | -106.0045
18.0425 | -106.0045
18.0462 | -106.0046
18.0499 | -106.0042
18.0535 | -106.0049
18.0572 | -106.007
18.0608 | -106.0102
18.0645 | -106.0128
18.0682 | -106.0171
18.0718 | -106.0205
18.0755 | -106.0195
18.0792 | -106.0178
18.0828 | -106.019
18.0865 | -106.021
18.0902 | -106.0207
18.0938 | -106.0206
18.0975 | -106.0166
18.1012 | -106.0119
18.1048 | -106.008
18.1085 | -106.0059
18.1122 | -106.0024
18.1158 | -105.9993
18.1195 | -105.996
18.1232 | -105.9899
18.1268 | -105.9867
18.1305 | -105.9813
18.1342 | -105.9732
18.1378 | -105.9664
18.1415 | -105.9576
18.1452 | -105.9487
18.1488 | -105.9417
18.1525 | -105.9316
18.1562 | -105.9223
18.1598 | -105.9134
18.1635 | -105.9024
18.1672 | -105.8931
18.1708 | -105.8862
18.1745 | -105.8784
18.1782 | -105.8727
18.1818 | -105.8663
18.1855 | -105.8598
18.1892 | -105.8537
18.1928 | -105.8496
18.1965 | -105.8443
18.2001 | -105.8412
18.2038 | -105.8377
18.2075 | -105.835
18.2111 | -105.8331
18.2148 | -105.8309
18.2185 | -105.8305
18.2221 | -105.8322
18.2258 | -105.8348
18.2295 | -105.8372
18.2331 | -105.8373
18.2368 | -105.8372
18.2405 | -105.8399
18.2441 | -105.8445
18.2478 | -105.8502
18.2515 | -105.8535
18.2551 | -105.8543
18.2588 | -105.859
18.2625 | -105.8658
18.2661 | -105.8742
18.2698 | -105.8809
18.2735 | -105.8887
18.2771 | -105.8959
18.2808 | -105.9023
18.2845 | -105.9112
18.2881 | -105.9187
18.2918 | -105.9265
18.2955 | -105.9329
18.2991 | -105.9409
18.3028 | -105.9472
18.3065 | -105.9523
18.3101 | -105.9595
18.3138 | -105.9643
18.3174 | -105.9659
18.3211 | -105.9671
18.3248 | -105.9691
18.3284 | -105.9707
18.3321 | -105.9735
18.3358 | -105.9766
18.3394 | -105.9802
18.3431 | -105.9822
18.3468 | -105.9839
18.3504 | -105.9861
18.3541 | -105.9856
18.3578 | -105.99
18.3614 | -105.995
18.3651 | -105.9974
18.3688 | -105.9957
18.3724 | -105.9957
18.3761 | -105.9959
18.3798 | -105.9954
18.3834 | -105.9937
18.3871 | -105.9893
18.3908 | -105.9899
18.3944 | -105.9905
18.3981 | -105.9887
18.4018 | -105.9856
18.4054 | -105.9792
18.4091 | -105.9728
18.4128 | -105.9664
18.4164 | -105.9592
18.4201 | -105.9493
18.4238 | -105.9402
18.4274 | -105.9313
18.4311 | -105.9204
18.4348 | -105.9113
18.4384 | -105.901
18.4421 | -105.8907
18.4457 | -105.8803
18.4494 | -105.8668
18.4531 | -105.8516
18.4567 | -105.8385
18.4604 | -105.8248
18.4641 | -105.8103
18.4677 | -105.7952
18.4714 | -105.7828
18.4751 | -105.772
18.4787 | -105.759
18.4824 | -105.7471
18.4861 | -105.7342
18.4897 | -105.7257
18.4934 | -105.719
18.4971 | -105.7121
18.5007 | -105.7031
18.5044 | -105.6935
18.5081 | -105.6928
18.5117 | -105.6871
18.5154 | -105.6833
18.5191 | -105.6832
18.5227 | -105.6817
18.5264 | -105.679
18.5301 | -105.6785
18.5337 | -105.6792
18.5374 | -105.6792
18.5411 | -105.6798
18.5447 | -105.6808
18.5484 | -105.6813
18.5521 | -105.6854
18.5557 | -105.6904
18.5594 | -105.6956
18.5631 | -105.7008
18.5667 | -105.7083
18.5704 | -105.7174
18.574 | -105.7254
18.5777 | -105.7325
18.5814 | -105.7416
18.585 | -105.7496
18.5887 | -105.7553
18.5924 | -105.7612
18.596 | -105.7662
18.5997 | -105.7736
18.6034 | -105.7814
18.607 | -105.788
18.6107 | -105.7961
18.6144 | -105.8035
18.618 | -105.8118
18.6217 | -105.817
18.6254 | -105.8228
18.629 | -105.8299
18.6327 | -105.8351
18.6364 | -105.8391
18.64 | -105.8435
18.6437 | -105.85
18.6474 | -105.8566
18.651 | -105.8607
18.6547 | -105.8613
18.6584 | -105.8624
18.662 | -105.8661
18.6657 | -105.8709
18.6694 | -105.8759
18.673 | -105.8788
18.6767 | -105.8812
18.6804 | -105.882
18.684 | -105.8861
18.6877 | -105.8894
18.6913 | -105.8904
18.695 | -105.8914
18.6987 | -105.8912
18.7023 | -105.8916
18.706 | -105.8912
18.7097 | -105.893
18.7133 | -105.8945
18.717 | -105.8959
18.7207 | -105.8974
18.7243 | -105.9009
18.728 | -105.9032
18.7317 | -105.9046
18.7353 | -105.9071
18.739 | -105.9082
18.7427 | -105.9076
18.7463 | -105.9084
18.75 | -105.9129
18.7537 | -105.9142
18.7573 | -105.9155
18.761 | -105.9173
18.7647 | -105.9216
18.7683 | -105.9275
18.772 | -105.9313
18.7757 | -105.9345
18.7793 | -105.9373
18.783 | -105.9406
18.7867 | -105.9445
18.7903 | -105.9489
18.794 | -105.9547
18.7977 | -105.9585
18.8013 | -105.9629
18.805 | -105.9686
18.8087 | -105.9738
18.8123 | -105.9791
18.816 | -105.9866
18.8196 | -105.9928
18.8233 | -105.9985
18.827 | -106.0043
18.8306 | -106.0092
18.8343 | -106.0121
18.838 | -106.016
18.8416 | -106.0222
18.8453 | -106.0297
18.849 | -106.0375
18.8526 | -106.0473
18.8563 | -106.0591
18.86 | -106.0707
18.8636 | -106.082
18.8673 | -106.0932
18.871 | -106.1027
18.8746 | -106.1126
18.8783 | -106.1198
18.882 | -106.1289
18.8856 | -106.1386
18.8893 | -106.1492
18.893 | -106.1584
18.8966 | -106.1661
18.9003 | -106.1745
18.904 | -106.1825
18.9076 | -106.188
18.9113 | -106.1912
18.915 | -106.1935
18.9186 | -106.1971
18.9223 | -106.2012
18.926 | -106.2038
18.9296 | -106.2041
18.9333 | -106.2043
18.9369 | -106.204
18.9406 | -106.205
18.9443 | -106.2058
18.9479 | -106.2049
18.9516 | -106.2013
18.9553 | -106.1996
18.9589 | -106.1959
18.9626 | -106.192
18.9663 | -106.1896
18.9699 | -106.1859
18.9736 | -106.1832
18.9773 | -106.1785
18.9809 | -106.1761
18.9846 | -106.1726
18.9883 | -106.1666
18.9919 | -106.1588
18.9956 | -106.1551
18.9993 | -106.1526
19.0029 | -106.1479
19.0066 | -106.1442
19.0103 | -106.1437
19.0139 | -106.1397
19.0176 | -106.1343
19.0213 | -106.1298
19.0249 | -106.1255
19.0286 | -106.1214
19.0323 | -106.1166
19.0359 | -106.114
19.0396 | -106.1066
19.0433 | -106.0996
19.0469 | -106.094
19.0506 | -106.0884
19.0543 | -106.0832
19.0579 | -106.0788
19.0616 | -106.0755
19.0652 | -106.0724
19.0689 | -106.0695
19.0726 | -106.0664
19.0762 | -106.0632
19.0799 | -106.0583
19.0836 | -106.0543
19.0872 | -106.051
19.0909 | -106.0469
19.0946 | -106.0431
19.0982 | -106.0421
19.1019 | -106.0413
19.1056 | -106.04
19.1092 | -106.0369
19.1129 | -106.0327
19.1166 | -106.0291
19.1202 | -106.0281
19.1239 | -106.0263
19.1276 | -106.0209
19.1312 | -106.0164
19.1349 | -106.014
19.1386 | -106.0104
19.1422 | -106.0079
19.1459 | -106.0051
19.1496 | -106.0026
19.1532 | -105.9989
19.1569 | -105.9969
19.1606 | -105.9952
19.1642 | -105.993
19.1679 | -105.9905
19.1716 | -105.9864
19.1752 | -105.982
19.1789 | -105.9782
19.1826 | -105.975
19.1862 | -105.9715
19.1899 | -105.9696
19.1935 | -105.9689
19.1972 | -105.9671
19.2009 | -105.9618
19.2045 | -105.9557
19.2082 | -105.9539
19.2119 | -105.9491
19.2155 | -105.9435
19.2192 | -105.9397
19.2229 | -105.9356
19.2265 | -105.9312
19.2302 | -105.9288
19.2339 | -105.9261
19.2375 | -105.9211
19.2412 | -105.9169
19.2449 | -105.9122
19.2485 | -105.9048
19.2522 | -105.8964
19.2559 | -105.888
19.2595 | -105.8821
19.2632 | -105.8758
19.2669 | -105.867
19.2705 | -105.8568
19.2742 | -105.8466
19.2779 | -105.8377
19.2815 | -105.8283
19.2852 | -105.8193
19.2889 | -105.8092
19.2925 | -105.7974
19.2962 | -105.7874
19.2999 | -105.778
19.3035 | -105.7694
19.3072 | -105.7603
19.3108 | -105.7505
19.3145 | -105.743
19.3182 | -105.7359
19.3218 | -105.7257
19.3255 | -105.7182
19.3292 | -105.7109
19.3328 | -105.7024
19.3365 | -105.6947
19.3402 | -105.6875
19.3438 | -105.6806
19.3475 | -105.6744
19.3512 | -105.6691
19.3548 | -105.6639
19.3585 | -105.6593
19.3622 | -105.6524
19.3658 | -105.648
19.3695 | -105.6452
19.3732 | -105.6417
19.3768 | -105.6383
19.3805 | -105.6378
19.3842 | -105.6371
19.3878 | -105.635
19.3915 | -105.6336
19.3952 | -105.6293
19.3988 | -105.6252
19.4025 | -105.6245
19.4062 | -105.6254
19.4098 | -105.6262
19.4135 | -105.6259
19.4172 | -105.6262
19.4208 | -105.6277
19.4245 | -105.6294
19.4282 | -105.6296
19.4318 | -105.6294
19.4355 | -105.6299
19.4392 | -105.6321
19.4428 | -105.635
19.4465 | -105.6375
19.4501 | -105.6388
19.4538 | -105.6414
19.4575 | -105.6426
19.4611 | -105.6438
19.4648 | -105.6452
19.4685 | -105.6451
19.4721 | -105.6455
19.4758 | -105.6494
19.4795 | -105.6544
19.4831 | -105.656
19.4868 | -105.6574
19.4905 | -105.6575
19.4941 | -105.6564
19.4978 | -105.6556
19.5015 | -105.6531
19.5051 | -105.6502
19.5088 | -105.6469
19.5125 | -105.6441
19.5161 | -105.6416
19.5198 | -105.6395
19.5235 | -105.6382
19.5271 | -105.6356
19.5308 | -105.6318
19.5345 | -105.6234
19.5381 | -105.6148
19.5418 | -105.6057
19.5455 | -105.5956
19.5491 | -105.5861
19.5528 | -105.5772
19.5565 | -105.5685
19.5601 | -105.5597
19.5638 | -105.545
19.5674 | -105.534
19.5711 | -105.5226
19.5748 | -105.5104
19.5784 | -105.4971
19.5821 | -105.4835
19.5858 | -105.4703
19.5894 | -105.4601
19.5931 | -105.4484
19.5968 | -105.4342
19.6004 | -105.4214
19.6041 | -105.4087
19.6078 | -105.3942
19.6114 | -105.3819
19.6151 | -105.3696
19.6188 | -105.3572
19.6224 | -105.346
19.6261 | -105.3355
19.6298 | -105.3273
19.6334 | -105.3197
19.6371 | -105.3122
19.6408 | -105.3069
19.6444 | -105.3003
19.6481 | -105.2934
19.6518 | -105.288
19.6554 | -105.2851
19.6591 | -105.2819
19.6628 | -105.2791
19.6664 | -105.2778
19.6701 | -105.2775
19.6738 | -105.2803
19.6774 | -105.2847
19.6811 | -105.2884
19.6848 | -105.2916
19.6884 | -105.2927
19.6921 | -105.2945
19.6957 | -105.2959
19.6994 | -105.2991
19.7031 | -105.3037
19.7067 | -105.3109
19.7104 | -105.3173
19.7141 | -105.3243
19.7177 | -105.3312
19.7214 | -105.3377
19.7251 | -105.3452
19.7287 | -105.3536
19.7324 | -105.3618
19.7361 | -105.3701
19.7397 | -105.3772
19.7434 | -105.3838
19.7471 | -105.391
19.7507 | -105.3981
19.7544 | -105.4031
19.7581 | -105.4068
19.7617 | -105.4116
19.7654 | -105.419
19.7691 | -105.4268
19.7727 | -105.4354
19.7764 | -105.4424
19.7801 | -105.4492
19.7837 | -105.4534
19.7874 | -105.4563
19.7911 | -105.4593
19.7947 | -105.4615
19.7984 | -105.4611
19.8021 | -105.4617
19.8057 | -105.4658
19.8094 | -105.4694
19.8131 | -105.4708
19.8167 | -105.4737
19.8204 | -105.4765
19.824 | -105.4776
19.8277 | -105.4767
19.8314 | -105.4751
19.835 | -105.4738
19.8387 | -105.4733
19.8424 | -105.4731
19.846 | -105.4713
19.8497 | -105.4703
19.8534 | -105.4712
19.857 | -105.4715
19.8607 | -105.4726
19.8644 | -105.4731
19.868 | -105.4725
19.8717 | -105.4703
19.8754 | -105.4697
19.879 | -105.4704
19.8827 | -105.4714
19.8864 | -105.4709
19.89 | -105.471
19.8937 | -105.4696
19.8974 | -105.4694
19.901 | -105.468
19.9047 | -105.4654
19.9084 | -105.4622
19.912 | -105.4575
19.9157 | -105.4555
19.9194 | -105.4542
19.923 | -105.4529
19.9267 | -105.4525
19.9304 | -105.4529
19.934 | -105.4513
19.9377 | -105.4493
19.9413 | -105.4484
19.945 | -105.4493
19.9487 | -105.4496
19.9523 | -105.4483
19.956 | -105.4466
19.9597 | -105.4445
19.9633 | -105.4435
19.967 | -105.4434
19.9707 | -105.4477
19.9743 | -105.4526
19.978 | -105.4566
19.9817 | -105.46
19.9853 | -105.4617
19.989 | -105.4621
19.9927 | -105.4619
19.9963 | -105.4605
20 | -105.462
20.0037 | -105.4646
20.0073 | -105.4685
20.011 | -105.4715
20.0147 | -105.4747
20.0183 | -105.4796
20.022 | -105.4866
20.0257 | -105.4925
20.0293 | -105.4969
20.033 | -105.5022
20.0367 | -105.5106
20.0403 | -105.5209
20.044 | -105.531
20.0477 | -105.5382
20.0513 | -105.5463
20.055 | -105.5543
20.0587 | -105.5636
20.0623 | -105.572
20.066 | -105.5819
20.0696 | -105.5916
20.0733 | -105.6019
20.077 | -105.6112
20.0806 | -105.6198
20.0843 | -105.6293
20.088 | -105.6377
20.0916 | -105.6444
20.0953 | -105.6494
20.099 | -105.6567
20.1026 | -105.6629
20.1063 | -105.6669
20.11 | -105.6707
20.1136 | -105.675
20.1173 | -105.6773
20.121 | -105.6792
20.1246 | -105.6794
20.1283 | -105.6807
20.132 | -105.6826
20.1356 | -105.6821
20.1393 | -105.6798
20.143 | -105.6781
20.1466 | -105.6742
20.1503 | -105.6678
20.154 | -105.6619
20.1576 | -105.6555
20.1613 | -105.6498
20.165 | -105.6429
20.1686 | -105.6342
20.1723 | -105.626
20.176 | -105.6197
20.1796 | -105.6128
20.1833 | -105.6043
20.1869 | -105.5976
20.1906 | -105.5905
20.1943 | -105.5849
20.1979 | -105.577
20.2016 | -105.5679
20.2053 | -105.5579
20.2089 | -105.548
20.2126 | -105.5405
20.2163 | -105.5318
20.2199 | -105.5228
20.2236 | -105.5171
20.2273 | -105.5118
20.2309 | -105.5058
20.2346 | -105.5004
20.2383 | -105.4965
20.2419 | -105.491
20.2456 | -105.484
20.2493 | -105.4768
20.2529 | -105.4725
20.2566 | -105.4668
20.2603 | -105.4616
20.2639 | -105.4564
20.2676 | -105.4512
20.2713 | -105.4477
20.2749 | -105.4436
20.2786 | -105.4401
20.2823 | -105.4369
20.2859 | -105.4306
20.2896 | -105.4245
20.2933 | -105.4188
20.2969 | -105.4147
20.3006 | -105.4079
20.3043 | -105.3984
20.3079 | -105.3882
20.3116 | -105.3803
20.3152 | -105.373
20.3189 | -105.3633
20.3226 | -105.3534
20.3262 | -105.3426
20.3299 | -105.3312
20.3336 | -105.3211
20.3372 | -105.3122
20.3409 | -105.3034
20.3446 | -105.2932
20.3482 | -105.2838
20.3519 | -105.2746
20.3556 | -105.2664
20.3592 | -105.2576
20.3629 | -105.2487
20.3666 | -105.2416
20.3702 | -105.2339
20.3739 | -105.2236
20.3776 | -105.2167
20.3812 | -105.2109
20.3849 | -105.2039
20.3886 | -105.2
20.3922 | -105.1963
20.3959 | -105.1958
20.3996 | -105.1972
20.4032 | -105.1974
20.4069 | -105.1984
20.4106 | -105.1987
20.4142 | -105.1996
20.4179 | -105.201
20.4216 | -105.2024
20.4252 | -105.2055
20.4289 | -105.212
20.4326 | -105.2192
20.4362 | -105.2281
20.4399 | -105.2393
20.4435 | -105.251
20.4472 | -105.2627
20.4509 | -105.2738
20.4545 | -105.2851
20.4582 | -105.2975
20.4619 | -105.3088
20.4655 | -105.3193
20.4692 | -105.3312
20.4729 | -105.3416
20.4765 | -105.3522
20.4802 | -105.3643
20.4839 | -105.3743
20.4875 | -105.3858
20.4912 | -105.3965
20.4949 | -105.4064
20.4985 | -105.416
20.5022 | -105.4235
20.5059 | -105.4323
20.5095 | -105.4413
20.5132 | -105.4475
20.5169 | -105.4534
20.5205 | -105.4591
20.5242 | -105.4622
20.5279 | -105.4654
20.5315 | -105.47
20.5352 | -105.4752
20.5389 | -105.4771
20.5425 | -105.4755
20.5462 | -105.4748
20.5499 | -105.4732
20.5535 | -105.4711
20.5572 | -105.4674
20.5608 | -105.4626
20.5645 | -105.46
20.5682 | -105.4567
20.5718 | -105.452
20.5755 | -105.4468
20.5792 | -105.4425
20.5828 | -105.4366
20.5865 | -105.4303
20.5902 | -105.4276
20.5938 | -105.4249
20.5975 | -105.4217
20.6012 | -105.4159
20.6048 | -105.4114
20.6085 | -105.4096
20.6122 | -105.4058
20.6158 | -105.4002
20.6195 | -105.3952
20.6232 | -105.3938
20.6268 | -105.3903
20.6305 | -105.3872
20.6342 | -105.3842
20.6378 | -105.3826
20.6415 | -105.3808
20.6452 | -105.3803
20.6488 | -105.3805
20.6525 | -105.3796
20.6562 | -105.3799
20.6598 | -105.3803
20.6635 | -105.3791
20.6672 | -105.3809
20.6708 | -105.3823
20.6745 | -105.3845
20.6782 | -105.3885
20.6818 | -105.3899
20.6855 | -105.3911
20.6892 | -105.3919
20.6928 | -105.3922
20.6965 | -105.3951
20.7001 | -105.3984
20.7038 | -105.4024
20.7075 | -105.4065
20.7111 | -105.4101
20.7148 | -105.4139
20.7185 | -105.417
20.7221 | -105.4166
20.7258 | -105.4137
20.7295 | -105.4109
20.7331 | -105.4117
20.7368 | -105.4131
20.7405 | -105.4145
20.7441 | -105.4142
20.7478 | -105.4149
20.7515 | -105.4142
20.7551 | -105.413
20.7588 | -105.4134
20.7625 | -105.4132
20.7661 | -105.4132
20.7698 | -105.411
20.7735 | -105.408
20.7771 | -105.4053
20.7808 | -105.4012
20.7845 | -105.397
20.7881 | -105.3933
20.7918 | -105.3896
20.7955 | -105.3843
20.7991 | -105.3779
20.8028 | -105.3728
20.8065 | -105.3678
20.8101 | -105.3632
20.8138 | -105.3587
20.8174 | -105.3551
20.8211 | -105.3504
20.8248 | -105.3425
20.8284 | -105.3362
20.8321 | -105.3331
20.8358 | -105.33
20.8394 | -105.326
20.8431 | -105.3222
20.8468 | -105.3179
20.8504 | -105.3128
20.8541 | -105.3086
20.8578 | -105.304
20.8614 | -105.2999
20.8651 | -105.2968
20.8688 | -105.2944
20.8724 | -105.2915
20.8761 | -105.2896
20.8798 | -105.2883
20.8834 | -105.2872
20.8871 | -105.2883
20.8908 | -105.2906
20.8944 | -105.2939
20.8981 | -105.2957
20.9018 | -105.2988
20.9054 | -105.3027
20.9091 | -105.3071
20.9128 | -105.3116
20.9164 | -105.3156
20.9201 | -105.3203
20.9238 | -105.327
20.9274 | -105.3334
20.9311 | -105.3403
20.9348 | -105.3477
20.9384 | -105.3548
20.9421 | -105.361
20.9457 | -105.3684
20.9494 | -105.3774
20.9531 | -105.3859
20.9567 | -105.3943
20.9604 | -105.4036
20.9641 | -105.4113
20.9677 | -105.4201
20.9714 | -105.4268
20.9751 | -105.4347
20.9787 | -105.4456
20.9824 | -105.4568
20.9861 | -105.4676
20.9897 | -105.4788
20.9934 | -105.4874
20.9971 | -105.4949
21.0007 | -105.5031
21.0044 | -105.5123
21.0081 | -105.5212
21.0117 | -105.5304
21.0154 | -105.5395
21.0191 | -105.5496
21.0227 | -105.5601
21.0264 | -105.5708
21.0301 | -105.5759
21.0337 | -105.5796
21.0374 | -105.5836
21.0411 | -105.5871
21.0447 | -105.59
21.0484 | -105.5951
21.0521 | -105.5997
21.0557 | -105.5999
21.0594 | -105.5988
21.0631 | -105.5978
21.0667 | -105.596
21.0704 | -105.5923
21.074 | -105.5875
21.0777 | -105.5844
21.0814 | -105.5784
21.085 | -105.5712
21.0887 | -105.5644
21.0924 | -105.5585
21.096 | -105.5508
21.0997 | -105.5418
21.1034 | -105.5317
21.107 | -105.5205
21.1107 | -105.5086
21.1144 | -105.4972
21.118 | -105.4845
21.1217 | -105.4747
21.1254 | -105.4628
21.129 | -105.4486
21.1327 | -105.4333
21.1364 | -105.4187
21.14 | -105.4058
21.1437 | -105.3938
21.1474 | -105.3815
21.151 | -105.3675
21.1547 | -105.3551
21.1584 | -105.3434
21.162 | -105.3318
21.1657 | -105.3207
21.1694 | -105.3092
21.173 | -105.2983
21.1767 | -105.2885
21.1804 | -105.2785
21.184 | -105.27
21.1877 | -105.2629
21.1913 | -105.255
21.195 | -105.2478
21.1987 | -105.2431
21.2023 | -105.2374
21.206 | -105.2329
21.2097 | -105.2325
21.2133 | -105.2331
21.217 | -105.234
21.2207 | -105.2345
21.2243 | -105.2363
21.228 | -105.2394
21.2317 | -105.2423
21.2353 | -105.2475
21.239 | -105.2549
21.2427 | -105.2608
21.2463 | -105.2649
21.25 | -105.2723
21.2537 | -105.2799
21.2573 | -105.2882
21.261 | -105.2965
21.2647 | -105.3067
21.2683 | -105.3172
21.272 | -105.3284
21.2757 | -105.3397
21.2793 | -105.3505
21.283 | -105.3603
21.2867 | -105.3719
21.2903 | -105.3823
21.294 | -105.3931
21.2977 | -105.4053
21.3013 | -105.4184
21.305 | -105.4304
21.3087 | -105.4416
21.3123 | -105.4514
21.316 | -105.4594
21.3196 | -105.4672
21.3233 | -105.475
21.327 | -105.482
21.3306 | -105.4907
21.3343 | -105.4982
21.338 | -105.5054
21.3416 | -105.5119
21.3453 | -105.5152
21.349 | -105.5182
21.3526 | -105.5209
21.3563 | -105.5242
21.36 | -105.5277
21.3636 | -105.5308
21.3673 | -105.5342
21.371 | -105.5368
21.3746 | -105.5376
21.3783 | -105.5361
21.382 | -105.5363
21.3856 | -105.5362
21.3893 | -105.5329
21.393 | -105.5279
21.3966 | -105.5244
21.4003 | -105.5216
21.404 | -105.5184
21.4076 | -105.5126
21.4113 | -105.5073
21.415 | -105.5028
21.4186 | -105.4997
21.4223 | -105.4934
21.426 | -105.4881
21.4296 | -105.4849
21.4333 | -105.4822
21.4369 | -105.4781
21.4406 | -105.4727
21.4443 | -105.4671
21.4479 | -105.4627
21.4516 | -105.4603
21.4553 | -105.4565
21.4589 | -105.4516
21.4626 | -105.4442
21.4663 | -105.4377
21.4699 | -105.4335
21.4736 | -105.4296
21.4773 | -105.4243
21.4809 | -105.4183
21.4846 | -105.4139
21.4883 | -105.4085
21.4919 | -105.402
21.4956 | -105.3976
21.4993 | -105.3919
21.5029 | -105.3852
21.5066 | -105.3805
21.5103 | -105.3771
21.5139 | -105.3718
21.5176 | -105.3653
21.5213 | -105.3593
21.5249 | -105.3512
21.5286 | -105.3416
21.5323 | -105.3339
21.5359 | -105.3279
21.5396 | -105.3206
21.5433 | -105.3124
21.5469 | -105.3059
21.5506 | -105.2995
21.5543 | -105.2916
21.5579 | -105.2859
21.5616 | -105.2802
21.5652 | -105.2741
21.5689 | -105.2646
21.5726 | -105.253
21.5762 | -105.2431
21.5799 | -105.2346
21.5836 | -105.2268
21.5872 | -105.2198
21.5909 | -105.2124
21.5946 | -105.2072
21.5982 | -105.2021
21.6019 | -105.1954
21.6056 | -105.1876
21.6092 | -105.1811
21.6129 | -105.1763
21.6166 | -105.1723
21.6202 | -105.1697
21.6239 | -105.166
21.6276 | -105.1627
21.6312 | -105.1613
21.6349 | -105.1594
21.6386 | -105.1587
21.6422 | -105.1592
21.6459 | -105.1593
21.6496 | -105.1612
21.6532 | -105.1616
21.6569 | -105.1622
21.6606 | -105.1648
21.6642 | -105.1663
21.6679 | -105.167
21.6716 | -105.1693
21.6752 | -105.1717
21.6789 | -105.1728
21.6826 | -105.1738
21.6862 | -105.1755
21.6899 | -105.1773
21.6935 | -105.1783
21.6972 | -105.1793
21.7009 | -105.1808
21.7045 | -105.1818
21.7082 | -105.1821
21.7119 | -105.183
21.7155 | -105.1857
21.7192 | -105.1881
21.7229 | -105.1897
21.7265 | -105.1904
21.7302 | -105.1931
21.7339 | -105.1946
21.7375 | -105.1938
21.7412 | -105.1909
21.7449 | -105.1895
21.7485 | -105.1895
21.7522 | -105.1916
21.7559 | -105.1933
21.7595 | -105.1931
21.7632 | -105.195
21.7669 | -105.1945
21.7705 | -105.1955
21.7742 | -105.1979
21.7779 | -105.1993
21.7815 | -105.2004
21.7852 | -105.2029
21.7889 | -105.2061
21.7925 | -105.2072
21.7962 | -105.2083
21.7999 | -105.2095
21.8035 | -105.2106
21.8072 | -105.2131
21.8108 | -105.2164
21.8145 | -105.2181
21.8182 | -105.2191
21.8218 | -105.2208
21.8255 | -105.2248
21.8292 | -105.2282
21.8328 | -105.2331
21.8365 | -105.2362
21.8402 | -105.2383
21.8438 | -105.2396
21.8475 | -105.2423
21.8512 | -105.2456
21.8548 | -105.2486
21.8585 | -105.252
21.8622 | -105.2513
21.8658 | -105.2496
21.8695 | -105.2509
21.8732 | -105.2535
21.8768 | -105.2584
21.8805 | -105.2627
21.8842 | -105.265
21.8878 | -105.2656
21.8915 | -105.2656
21.8952 | -105.2665
21.8988 | -105.2679
21.9025 | -105.2682
21.9062 | -105.269
21.9098 | -105.2699
21.9135 | -105.2715
21.9172 | -105.2722
21.9208 | -105.2722
21.9245 | -105.2729
21.9282 | -105.2751
21.9318 | -105.2772
21.9355 | -105.2783
21.9392 | -105.2785
21.9428 | -105.2788
21.9465 | -105.2788
21.9501 | -105.277
21.9538 | -105.2781
21.9575 | -105.2798
21.9611 | -105.2801
21.9648 | -105.2803
21.9685 | -105.2789
21.9721 | -105.2801
21.9758 | -105.2823
21.9795 | -105.2839
21.9831 | -105.2877
21.9868 | -105.2928
21.9905 | -105.2965
21.9941 | -105.2987
21.9978 | -105.3014
22.0015 | -105.305
22.0051 | -105.3096
22.0088 | -105.3143
22.0125 | -105.3184
22.0161 | -105.3232
22.0198 | -105.3271
22.0235 | -105.3295
22.0271 | -105.3318
22.0308 | -105.3356
22.0345 | -105.3365
22.0381 | -105.337
22.0418 | -105.339
22.0455 | -105.341
22.0491 | -105.3426
22.0528 | -105.345
22.0565 | -105.3453
22.0601 | -105.3446
22.0638 | -105.3405
22.0674 | -105.3351
22.0711 | -105.3302
22.0748 | -105.3251
22.0784 | -105.3221
22.0821 | -105.3187
22.0858 | -105.3114
22.0894 | -105.3015
22.0931 | -105.292
22.0968 | -105.2824
22.1004 | -105.2734
22.1041 | -105.2649
22.1078 | -105.2561
22.1114 | -105.2484
22.1151 | -105.2398
22.1188 | -105.2302
22.1224 | -105.219
22.1261 | -105.207
22.1298 | -105.1936
22.1334 | -105.1807
22.1371 | -105.1669
22.1408 | -105.1556
22.1444 | -105.1455
22.1481 | -105.136
22.1518 | -105.1253
22.1554 | -105.1148
22.1591 | -105.1051
22.1628 | -105.0951
22.1664 | -105.0862
22.1701 | -105.0773
22.1738 | -105.0722
22.1774 | -105.0682
22.1811 | -105.0607
22.1848 | -105.0532
22.1884 | -105.0467
22.1921 | -105.0435
22.1957 | -105.0399
22.1994 | -105.035
22.2031 | -105.0309
22.2067 | -105.0283
22.2104 | -105.0322
22.2141 | -105.0302
22.2177 | -105.029
22.2214 | -105.0299
22.2251 | -105.0327
22.2287 | -105.0343
22.2324 | -105.0364
22.2361 | -105.039
22.2397 | -105.0425
22.2434 | -105.0454
22.2471 | -105.0481
22.2507 | -105.0537
22.2544 | -105.0582
22.2581 | -105.0618
22.2617 | -105.0657
22.2654 | -105.071
22.2691 | -105.0765
22.2727 | -105.083
22.2764 | -105.0913
22.2801 | -105.0975
22.2837 | -105.1038
22.2874 | -105.1075
22.2911 | -105.1118
22.2947 | -105.117
22.2984 | -105.1207
22.3021 | -105.1243
22.3057 | -105.1288
22.3094 | -105.1324
22.3131 | -105.1353
22.3167 | -105.1361
22.3204 | -105.138
22.324 | -105.1396
22.3277 | -105.1397
22.3314 | -105.1393
22.335 | -105.1389
22.3387 | -105.1371
22.3424 | -105.1324
22.346 | -105.1272
22.3497 | -105.1255
22.3534 | -105.1232
22.357 | -105.1185
22.3607 | -105.1124
22.3644 | -105.1079
22.368 | -105.1035
22.3717 | -105.1004
22.3754 | -105.0942
22.379 | -105.0871
22.3827 | -105.0819
22.3864 | -105.074
22.39 | -105.0657
22.3937 | -105.0577
22.3974 | -105.0472
22.401 | -105.0394
22.4047 | -105.0338
22.4084 | -105.0267
22.412 | -105.0198
22.4157 | -105.0099
22.4194 | -105.0007
22.423 | -104.9917
22.4267 | -104.9827
22.4304 | -104.9743
22.434 | -104.9672
22.4377 | -104.9616
22.4413 | -104.9574
22.445 | -104.9514
22.4487 | -104.9449
22.4523 | -104.9396
22.456 | -104.9343
22.4597 | -104.9285
22.4633 | -104.9248
22.467 | -104.9217
22.4707 | -104.9195
22.4743 | -104.9164
22.478 | -104.9147
22.4817 | -104.913
22.4853 | -104.9112
22.489 | -104.9081
22.4927 | -104.9066
22.4963 | -104.9077
22.5 | -104.9099
22.5037 | -104.9109
22.5073 | -104.9117
22.511 | -104.9153
22.5147 | -104.9203
22.5183 | -104.9255
22.522 | -104.9289
22.5257 | -104.9322
22.5293 | -104.9363
22.533 | -104.9408
22.5367 | -104.9464
22.5403 | -104.9514
22.544 | -104.9566
22.5477 | -104.9635
22.5513 | -104.9704
22.555 | -104.9759
22.5587 | -104.9817
22.5623 | -104.9862
22.566 | -104.9906
22.5696 | -104.9937
22.5733 | -104.9989
22.577 | -105.0065
22.5806 | -105.0137
22.5843 | -105.0204
22.588 | -105.0267
22.5916 | -105.0309
22.5953 | -105.0359
22.599 | -105.0435
22.6026 | -105.0498
22.6063 | -105.0557
22.61 | -105.0607
22.6136 | -105.0661
22.6173 | -105.0696
22.621 | -105.073
22.6246 | -105.0769
22.6283 | -105.0811
22.632 | -105.0863
22.6356 | -105.0912
22.6393 | -105.0942
22.643 | -105.0982
22.6466 | -105.1014
22.6503 | -105.1031
22.654 | -105.1032
22.6576 | -105.1054
22.6613 | -105.1086
22.665 | -105.1093
22.6686 | -105.1093
22.6723 | -105.1088
22.676 | -105.108
22.6796 | -105.1063
22.6833 | -105.1033
22.6869 | -105.0995
22.6906 | -105.0931
22.6943 | -105.0874
22.6979 | -105.0848
22.7016 | -105.0813
22.7053 | -105.0779
22.7089 | -105.0735
22.7126 | -105.0679
22.7163 | -105.0624
22.7199 | -105.0532
22.7236 | -105.0415
22.7273 | -105.0315
22.7309 | -105.0193
22.7346 | -105.0087
22.7383 | -104.9903
22.7419 | -104.9795
22.7456 | -104.9684
22.7493 | -104.9533
22.7529 | -104.9392
22.7566 | -104.9256
22.7603 | -104.9109
22.7639 | -104.8955
22.7676 | -104.8819
22.7713 | -104.8678
22.7749 | -104.8516
22.7786 | -104.8327
22.7823 | -104.8152
22.7859 | -104.7979
22.7896 | -104.7808
22.7933 | -104.7617
22.7969 | -104.743
22.8006 | -104.7237
22.8043 | -104.705
22.8079 | -104.6876
22.8116 | -104.6707
22.8152 | -104.6527
22.8189 | -104.6349
22.8226 | -104.6166
22.8262 | -104.5984
22.8299 | -104.5801
22.8336 | -104.5611
22.8372 | -104.5425
22.8409 | -104.5247
22.8446 | -104.5095
22.8482 | -104.4952
22.8519 | -104.4808
22.8556 | -104.4674
22.8592 | -104.4549
22.8629 | -104.4413
22.8666 | -104.4289
22.8702 | -104.4169
22.8739 | -104.4068
22.8776 | -104.4001
22.8812 | -104.3928
22.8849 | -104.3876
22.8886 | -104.3842
22.8922 | -104.381
22.8959 | -104.3781
22.8996 | -104.3736
22.9032 | -104.3701
22.9069 | -104.3675
22.9106 | -104.3662
22.9142 | -104.367
22.9179 | -104.3666
22.9216 | -104.3662
22.9252 | -104.3674
22.9289 | -104.3692
22.9326 | -104.3723
22.9362 | -104.3743
22.9399 | -104.3765
22.9435 | -104.3802
22.9472 | -104.3851
22.9509 | -104.3897
22.9545 | -104.392
22.9582 | -104.3957
22.9619 | -104.3998
22.9655 | -104.4028
22.9692 | -104.408
22.9729 | -104.4133
22.9765 | -104.4182
22.9802 | -104.4236
22.9839 | -104.4288
22.9875 | -104.4337
22.9912 | -104.4382
22.9949 | -104.4422
22.9985 | -104.4456
23.0022 | -104.4512
23.0059 | -104.4566
23.0095 | -104.4633
23.0132 | -104.4704
23.0169 | -104.4775
23.0205 | -104.4831
23.0242 | -104.4874
23.0279 | -104.4905
23.0315 | -104.4962
23.0352 | -104.5038
23.0389 | -104.5118
23.0425 | -104.5215
23.0462 | -104.5324
23.0499 | -104.5425
23.0535 | -104.554
23.0572 | -104.5645
23.0608 | -104.5754
23.0645 | -104.5863
23.0682 | -104.5982
23.0718 | -104.6097
23.0755 | -104.6234
23.0792 | -104.6344
23.0828 | -104.6452
23.0865 | -104.6571
23.0902 | -104.6697
23.0938 | -104.6822
23.0975 | -104.695
23.1012 | -104.7104
23.1048 | -104.7259
23.1085 | -104.7403
23.1122 | -104.7538
23.1158 | -104.7681
23.1195 | -104.785
23.1232 | -104.7992
23.1268 | -104.8121
23.1305 | -104.8238
23.1342 | -104.8347
23.1378 | -104.8477
23.1415 | -104.8591
23.1452 | -104.8688
23.1488 | -104.8801
23.1525 | -104.8925
23.1562 | -104.9043
23.1598 | -104.9157
23.1635 | -104.9257
23.1672 | -104.9342
23.1708 | -104.9408
23.1745 | -104.9464
23.1782 | -104.9517
23.1818 | -104.9553
23.1855 | -104.9579
23.1892 | -104.9613
23.1928 | -104.9641
23.1965 | -104.9658
23.2001 | -104.9652
23.2038 | -104.9638
23.2075 | -104.9598
23.2111 | -104.9563
23.2148 | -104.9526
23.2185 | -104.949
23.2221 | -104.9452
23.2258 | -104.9425
23.2295 | -104.9394
23.2331 | -104.9354
23.2368 | -104.93
23.2405 | -104.925
23.2441 | -104.9212
23.2478 | -104.9167
23.2515 | -104.9122
23.2551 | -104.9064
23.2588 | -104.9006
23.2625 | -104.8953
23.2661 | -104.8947
23.2698 | -104.8924
23.2735 | -104.8903
23.2771 | -104.8863
23.2808 | -104.8834
23.2845 | -104.8834
23.2881 | -104.8823
23.2918 | -104.8799
23.2955 | -104.8776
23.2991 | -104.8768
23.3028 | -104.8769
23.3065 | -104.8791
23.3101 | -104.8799
23.3138 | -104.8814
23.3174 | -104.8853
23.3211 | -104.8896
23.3248 | -104.8954
23.3284 | -104.9003
23.3321 | -104.9024
23.3358 | -104.9057
23.3394 | -104.9088
23.3431 | -104.9117
23.3468 | -104.9164
23.3504 | -104.9219
23.3541 | -104.9272
23.3578 | -104.9346
23.3614 | -104.938
23.3651 | -104.9401
23.3688 | -104.9455
23.3724 | -104.9511
23.3761 | -104.9572
23.3798 | -104.9631
23.3834 | -104.9681
23.3871 | -104.9729
23.3908 | -104.9771
23.3944 | -104.9817
23.3981 | -104.9853
23.4018 | -104.9896
23.4054 | -104.9948
23.4091 | -105.0025
23.4128 | -105.0053
23.4164 | -105.0079
23.4201 | -105.0118
23.4238 | -105.0171
23.4274 | -105.0212
23.4311 | -105.0228
23.4348 | -105.0247
23.4384 | -105.0271
23.4421 | -105.0295
23.4457 | -105.0314
23.4494 | -105.033
23.4531 | -105.0372
23.4567 | -105.0413
23.4604 | -105.0448
23.4641 | -105.0474
23.4677 | -105.0496
23.4714 | -105.0528
23.4751 | -105.0562
23.4787 | -105.0571
23.4824 | -105.0576
23.4861 | -105.06
23.4897 | -105.0631
23.4934 | -105.0678
23.4971 | -105.0727
23.5007 | -105.0768
23.5044 | -105.0796
23.5081 | -105.0833
23.5117 | -105.0888
23.5154 | -105.0933
23.5191 | -105.0987
23.5227 | -105.1025
23.5264 | -105.1078
23.5301 | -105.1129
23.5337 | -105.1177
23.5374 | -105.1263
23.5411 | -105.1339
23.5447 | -105.14
23.5484 | -105.1438
23.5521 | -105.1459
23.5557 | -105.15
23.5594 | -105.155
23.5631 | -105.1594
23.5667 | -105.1622
23.5704 | -105.1651
23.574 | -105.1684
23.5777 | -105.1722
23.5814 | -105.1772
23.585 | -105.1823
23.5887 | -105.1844
23.5924 | -105.188
23.596 | -105.1917
23.5997 | -105.193
23.6034 | -105.1958
23.607 | -105.1972
23.6107 | -105.1985
23.6144 | -105.2
23.618 | -105.2004
23.6217 | -105.2013
23.6254 | -105.2032
23.629 | -105.2052
23.6327 | -105.2049
23.6364 | -105.2043
23.64 | -105.2022
23.6437 | -105.2012
23.6474 | -105.1998
23.651 | -105.1975
23.6547 | -105.1945
23.6584 | -105.1922
23.662 | -105.1914
23.6657 | -105.1926
23.6694 | -105.1911
23.673 | -105.1887
23.6767 | -105.1878
23.6804 | -105.1865
23.684 | -105.184
23.6877 | -105.1819
23.6913 | -105.1811
23.695 | -105.1804
23.6987 | -105.178
23.7023 | -105.1742
23.706 | -105.1715
23.7097 | -105.1689
23.7133 | -105.1672
23.717 | -105.1674
23.7207 | -105.1701
23.7243 | -105.1734
23.728 | -105.1758
23.7317 | -105.1778
23.7353 | -105.1806
23.739 | -105.1843
23.7427 | -105.1866
23.7463 | -105.1927
23.75 | -105.1971
23.7537 | -105.2007
23.7573 | -105.2046
23.761 | -105.2091
23.7647 | -105.2145
23.7683 | -105.2214
23.772 | -105.2278
23.7757 | -105.2351
23.7793 | -105.2428
23.783 | -105.25
23.7867 | -105.2564
23.7903 | -105.2643
23.794 | -105.2726
23.7977 | -105.284
23.8013 | -105.2947
23.805 | -105.3026
23.8087 | -105.3108
23.8123 | -105.321
23.816 | -105.3306
23.8196 | -105.3406
23.8233 | -105.3506
23.827 | -105.3604
23.8306 | -105.3695
23.8343 | -105.3788
23.838 | -105.3896
23.8416 | -105.4029
23.8453 | -105.4148
23.849 | -105.4252
23.8526 | -105.4375
23.8563 | -105.4501
23.86 | -105.4618
23.8636 | -105.4733
23.8673 | -105.4862
23.871 | -105.4989
23.8746 | -105.5112
23.8783 | -105.524
23.882 | -105.5366
23.8856 | -105.5502
23.8893 | -105.5658
23.893 | -105.579
23.8966 | -105.5905
23.9003 | -105.6017
23.904 | -105.6132
23.9076 | -105.6235
23.9113 | -105.6345
23.915 | -105.646
23.9186 | -105.6547
23.9223 | -105.6646
23.926 | -105.6745
23.9296 | -105.6817
23.9333 | -105.687
23.9369 | -105.6938
23.9406 | -105.6993
23.9443 | -105.7054
23.9479 | -105.711
23.9516 | -105.7157
23.9553 | -105.7196
23.9589 | -105.7203
23.9626 | -105.7203
23.9663 | -105.7227
23.9699 | -105.7236
23.9736 | -105.7234
23.9773 | -105.7243
23.9809 | -105.7244
23.9846 | -105.7241
23.9883 | -105.7238
23.9919 | -105.7229
23.9956 | -105.7201
23.9993 | -105.7183
24.0029 | -105.7158
24.0066 | -105.7158
24.0103 | -105.7162
24.0139 | -105.7148
24.0176 | -105.712
24.0213 | -105.7099
24.0249 | -105.7081
24.0286 | -105.7057
24.0323 | -105.7051
24.0359 | -105.7048
24.0396 | -105.7046
24.0433 | -105.7044
24.0469 | -105.7045
24.0506 | -105.7053
24.0543 | -105.7056
24.0579 | -105.7044
24.0616 | -105.7051
24.0652 | -105.7059
24.0689 | -105.707
24.0726 | -105.7091
24.0762 | -105.7111
24.0799 | -105.7139
24.0836 | -105.7156
24.0872 | -105.7185
24.0909 | -105.7221
24.0946 | -105.7243
24.0982 | -105.726
24.1019 | -105.7277
24.1056 | -105.7294
24.1092 | -105.7308
24.1129 | -105.7312
24.1166 | -105.7319
24.1202 | -105.7343
24.1239 | -105.7375
24.1276 | -105.7381
24.1312 | -105.7397
24.1349 | -105.7406
24.1386 | -105.7408
24.1422 | -105.7394
24.1459 | -105.7367
24.1496 | -105.7349
24.1532 | -105.7332
24.1569 | -105.7346
24.1606 | -105.7363
24.1642 | -105.7347
24.1679 | -105.7333
24.1716 | -105.7332
24.1752 | -105.7324
24.1789 | -105.7319
24.1826 | -105.7295
24.1862 | -105.7279
24.1899 | -105.7276
24.1935 | -105.7266
24.1972 | -105.7251
24.2009 | -105.724
24.2045 | -105.7209
24.2082 | -105.7158
24.2119 | -105.7101
24.2155 | -105.7038
24.2192 | -105.6998
24.2229 | -105.6972
24.2265 | -105.6958
24.2302 | -105.6918
24.2339 | -105.6876
24.2375 | -105.6824
24.2412 | -105.678
24.2449 | -105.6752
24.2485 | -105.6722
24.2522 | -105.6662
24.2559 | -105.6605
24.2595 | -105.6561
24.2632 | -105.6502
24.2669 | -105.6454
24.2705 | -105.6387
24.2742 | -105.6336
24.2779 | -105.6263
24.2815 | -105.6159
24.2852 | -105.6051
24.2889 | -105.5981
24.2925 | -105.5869
24.2962 | -105.5729
24.2999 | -105.5589
24.3035 | -105.5453
24.3072 | -105.5288
24.3108 | -105.5155
24.3145 | -105.5037
24.3182 | -105.4909
24.3218 | -105.4771
24.3255 | -105.4592
24.3292 | -105.4459
24.3328 | -105.4306
24.3365 | -105.4143
24.3402 | -105.3999
24.3438 | -105.3848
24.3475 | -105.3697
24.3512 | -105.3535
24.3548 | -105.3364
24.3585 | -105.3199
24.3622 | -105.3033
24.3658 | -105.2878
24.3695 | -105.2723
24.3732 | -105.2552
24.3768 | -105.2369
24.3805 | -105.2203
24.3842 | -105.2014
24.3878 | -105.1839
24.3915 | -105.167
24.3952 | -105.152
24.3988 | -105.1385
24.4025 | -105.1255
24.4062 | -105.1121
24.4098 | -105.1016
24.4135 | -105.0915
24.4172 | -105.0801
24.4208 | -105.0697
24.4245 | -105.0616
24.4282 | -105.0572
24.4318 | -105.0532
24.4355 | -105.0492
24.4392 | -105.0485
24.4428 | -105.0497
24.4465 | -105.0512
24.4501 | -105.0536
24.4538 | -105.0583
24.4575 | -105.0651
24.4611 | -105.072
24.4648 | -105.0796
24.4685 | -105.09
24.4721 | -105.1031
24.4758 | -105.1175
24.4795 | -105.1317
24.4831 | -105.1478
24.4868 | -105.1656
24.4905 | -105.187
24.4941 | -105.2082
24.4978 | -105.2279
24.5015 | -105.2475
24.5051 | -105.2685
24.5088 | -105.2925
24.5125 | -105.3168
24.5161 | -105.3404
24.5198 | -105.3638
24.5235 | -105.3868
24.5271 | -105.4125
24.5308 | -105.4386
24.5345 | -105.462
24.5381 | -105.484
24.5418 | -105.5077
24.5455 | -105.5317
24.5491 | -105.5561
24.5528 | -105.5804
24.5565 | -105.6031
24.5601 | -105.6245
24.5638 | -105.6412
24.5674 | -105.6559
24.5711 | -105.6678
24.5748 | -105.6759
24.5784 | -105.6834
24.5821 | -105.691
24.5858 | -105.6941
24.5894 | -105.6887
24.5931 | -105.6853
24.5968 | -105.684
24.6004 | -105.6822
24.6041 | -105.678
24.6078 | -105.672
24.6114 | -105.6638
24.6151 | -105.6559
24.6188 | -105.6456
24.6224 | -105.6351
24.6261 | -105.6238
24.6298 | -105.6111
24.6334 | -105.5953
24.6371 | -105.5799
24.6408 | -105.563
24.6444 | -105.545
24.6481 | -105.5229
24.6518 | -105.4993
24.6554 | -105.4721
24.6591 | -105.4433
24.6628 | -105.4129
24.6664 | -105.3812
24.6701 | -105.3489
24.6738 | -105.3162
24.6774 | -105.2847
24.6811 | -105.2551
24.6848 | -105.2265
24.6884 | -105.1957
24.6921 | -105.1653
24.6957 | -105.1366
24.6994 | -105.1079
24.7031 | -105.0799
24.7067 | -105.0533
24.7104 | -105.0282
24.7141 | -105.0031
24.7177 | -104.9792
24.7214 | -104.9587
24.7251 | -104.9403
24.7287 | -104.9233
24.7324 | -104.9089
24.7361 | -104.8953
24.7397 | -104.8831
24.7434 | -104.8712
24.7471 | -104.8607
24.7507 | -104.8515
24.7544 | -104.8453
24.7581 | -104.8419
24.7617 | -104.8406
24.7654 | -104.8435
24.7691 | -104.8483
24.7727 | -104.8543
24.7764 | -104.8605
24.7801 | -104.8683
24.7837 | -104.8786
24.7874 | -104.8898
24.7911 | -104.9042
24.7947 | -104.9215
24.7984 | -104.9394
24.8021 | -104.9583
24.8057 | -104.9783
24.8094 | -105.0003
24.8131 | -105.0221
24.8167 | -105.0455
24.8204 | -105.0702
24.824 | -105.0935
24.8277 | -105.1191
24.8314 | -105.1452
24.835 | -105.1722
24.8387 | -105.1984
24.8424 | -105.2249
24.846 | -105.2529
24.8497 | -105.2799
24.8534 | -105.3018
24.857 | -105.3258
24.8607 | -105.3485
24.8644 | -105.3729
24.868 | -105.3975
24.8717 | -105.4219
24.8754 | -105.4445
24.879 | -105.4657
24.8827 | -105.487
24.8864 | -105.5082
24.89 | -105.5292
24.8937 | -105.5478
24.8974 | -105.5659
24.901 | -105.5837
24.9047 | -105.6015
24.9084 | -105.6188
24.912 | -105.6359
24.9157 | -105.6532
24.9194 | -105.6705
24.923 | -105.6873
24.9267 | -105.7051
24.9304 | -105.7222
24.934 | -105.7378
24.9377 | -105.7553
24.9413 | -105.7753
24.945 | -105.7892
24.9487 | -105.7993
24.9523 | -105.8062
24.956 | -105.8124
24.9597 | -105.82
24.9633 | -105.8294
24.967 | -105.8385
24.9707 | -105.8477
24.9743 | -105.8556
24.978 | -105.8643
24.9817 | -105.8716
24.9853 | -105.878
24.989 | -105.8864
24.9927 | -105.8965
24.9963 | -105.9069
25 | -105.9172
25.0037 | -105.9273
25.0073 | -105.9374
25.011 | -105.9457
25.0147 | -105.9547
25.0183 | -105.9645
25.022 | -105.9752
25.0257 | -105.9836
25.0293 | -105.9893
25.033 | -105.9949
25.0367 | -106.0003
25.0403 | -106.0056
25.044 | -106.0127
25.0477 | -106.0188
25.0513 | -106.0249
25.055 | -106.0313
25.0587 | -106.0371
25.0623 | -106.0423
25.066 | -106.0476
25.0696 | -106.055
25.0733 | -106.0603
25.077 | -106.0632
25.0806 | -106.065
25.0843 | -106.0662
25.088 | -106.068
25.0916 | -106.0697
25.0953 | -106.0714
25.099 | -106.073
25.1026 | -106.0775
25.1063 | -106.0804
25.11 | -106.0804
25.1136 | -106.079
25.1173 | -106.0772
25.121 | -106.0759
25.1246 | -106.0754
25.1283 | -106.0739
25.132 | -106.0735
25.1356 | -106.0719
25.1393 | -106.0689
25.143 | -106.0664
25.1466 | -106.0626
25.1503 | -106.0594
25.154 | -106.0567
25.1576 | -106.054
25.1613 | -106.051
25.165 | -106.0476
25.1686 | -106.0435
25.1723 | -106.0401
25.176 | -106.0375
25.1796 | -106.0333
25.1833 | -106.0279
25.1869 | -106.022
25.1906 | -106.0169
25.1943 | -106.0109
25.1979 | -106.0052
25.2016 | -106.0014
25.2053 | -105.9974
25.2089 | -105.9941
25.2126 | -105.9902
25.2163 | -105.9858
25.2199 | -105.9789
25.2236 | -105.9714
25.2273 | -105.965
25.2309 | -105.9598
25.2346 | -105.9552
25.2383 | -105.9495
25.2419 | -105.9455
25.2456 | -105.9406
25.2493 | -105.9373
25.2529 | -105.9336
25.2566 | -105.9291
25.2603 | -105.9256
25.2639 | -105.9232
25.2676 | -105.9184
25.2713 | -105.9125
25.2749 | -105.9082
25.2786 | -105.9053
25.2823 | -105.9041
25.2859 | -105.9008
25.2896 | -105.8967
25.2933 | -105.8918
25.2969 | -105.8888
25.3006 | -105.8845
25.3043 | -105.8806
25.3079 | -105.8777
25.3116 | -105.8747
25.3152 | -105.871
25.3189 | -105.8652
25.3226 | -105.8591
25.3262 | -105.8542
25.3299 | -105.8499
25.3336 | -105.8458
25.3372 | -105.8423
25.3409 | -105.8388
25.3446 | -105.8347
25.3482 | -105.8313
25.3519 | -105.8268
25.3556 | -105.8224
25.3592 | -105.8179
25.3629 | -105.8122
25.3666 | -105.8084
25.3702 | -105.8043
25.3739 | -105.8003
25.3776 | -105.7966
25.3812 | -105.7934
25.3849 | -105.7912
25.3886 | -105.7899
25.3922 | -105.7874
25.3959 | -105.7857
25.3996 | -105.7844
25.4032 | -105.7817
25.4069 | -105.7788
25.4106 | -105.7738
25.4142 | -105.7691
25.4179 | -105.7655
25.4216 | -105.7627
25.4252 | -105.761
25.4289 | -105.7574
25.4326 | -105.7541
25.4362 | -105.7535
25.4399 | -105.7531
25.4435 | -105.753
25.4472 | -105.7526
25.4509 | -105.7508
25.4545 | -105.7501
25.4582 | -105.7504
25.4619 | -105.7479
25.4655 | -105.7466
25.4692 | -105.7466
25.4729 | -105.7466
25.4765 | -105.7461
25.4802 | -105.745
25.4839 | -105.7438
25.4875 | -105.7436
25.4912 | -105.7431
25.4949 | -105.7411
25.4985 | -105.741
25.5022 | -105.7383
25.5059 | -105.7377
25.5095 | -105.7374
25.5132 | -105.7365
25.5169 | -105.7356
25.5205 | -105.7358
25.5242 | -105.7376
25.5279 | -105.7382
25.5315 | -105.7369
25.5352 | -105.7351
25.5389 | -105.7338
25.5425 | -105.7334
25.5462 | -105.7327
25.5499 | -105.7312
25.5535 | -105.7302
25.5572 | -105.7312
25.5608 | -105.7313
25.5645 | -105.7312
25.5682 | -105.7295
25.5718 | -105.7298
25.5755 | -105.7316
25.5792 | -105.7329
25.5828 | -105.7339
25.5865 | -105.7349
25.5902 | -105.7349
25.5938 | -105.7334
25.5975 | -105.7329
25.6012 | -105.7339
25.6048 | -105.7348
25.6085 | -105.734
25.6122 | -105.7328
25.6158 | -105.7321
25.6195 | -105.7315
25.6232 | -105.7296
25.6268 | -105.7278
25.6305 | -105.7289
25.6342 | -105.7286
25.6378 | -105.728
25.6415 | -105.727
25.6452 | -105.7268
25.6488 | -105.7283
25.6525 | -105.7301
25.6562 | -105.7302
25.6598 | -105.7316
25.6635 | -105.7333
25.6672 | -105.7356
25.6708 | -105.7369
25.6745 | -105.7371
25.6782 | -105.7362
25.6818 | -105.7355
25.6855 | -105.7335
25.6892 | -105.7324
25.6928 | -105.7327
25.6965 | -105.732
25.7001 | -105.7309
25.7038 | -105.7314
25.7075 | -105.7326
25.7111 | -105.7347
25.7148 | -105.736
25.7185 | -105.7362
25.7221 | -105.7368
25.7258 | -105.7366
25.7295 | -105.738
25.7331 | -105.7406
25.7368 | -105.7405
25.7405 | -105.7405
25.7441 | -105.7415
25.7478 | -105.7428
25.7515 | -105.7443
25.7551 | -105.7457
25.7588 | -105.7469
25.7625 | -105.7484
25.7661 | -105.7505
25.7698 | -105.7527
25.7735 | -105.7578
25.7771 | -105.7625
25.7808 | -105.7658
25.7845 | -105.7684
25.7881 | -105.7701
25.7918 | -105.7724
25.7955 | -105.7762
25.7991 | -105.78
25.8028 | -105.7825
25.8065 | -105.7855
25.8101 | -105.7884
25.8138 | -105.7931
25.8174 | -105.7981
25.8211 | -105.803
25.8248 | -105.8102
25.8284 | -105.816
25.8321 | -105.8215
25.8358 | -105.8272
25.8394 | -105.8331
25.8431 | -105.839
25.8468 | -105.8441
25.8504 | -105.8504
25.8541 | -105.8566
25.8578 | -105.8659
25.8614 | -105.8732
25.8651 | -105.8794
25.8688 | -105.8844
25.8724 | -105.8889
25.8761 | -105.8942
25.8798 | -105.9005
25.8834 | -105.9063
25.8871 | -105.9117
25.8908 | -105.9164
25.8944 | -105.9234
25.8981 | -105.9317
25.9018 | -105.9399
25.9054 | -105.9462
25.9091 | -105.9533
25.9128 | -105.9578
25.9164 | -105.9644
25.9201 | -105.9695
25.9238 | -105.9707
25.9274 | -105.974
25.9311 | -105.9778
25.9348 | -105.9821
25.9384 | -105.9855
25.9421 | -105.9887
25.9457 | -105.992
25.9494 | -105.9954
25.9531 | -105.9993
25.9567 | -106.0027
25.9604 | -106.0067
25.9641 | -106.0111
25.9677 | -106.0134
25.9714 | -106.0161
25.9751 | -106.0178
25.9787 | -106.0202
25.9824 | -106.0216
25.9861 | -106.0234
25.9897 | -106.0256
25.9934 | -106.0295
25.9971 | -106.0315
26.0007 | -106.0328
26.0044 | -106.0335
26.0081 | -106.0336
26.0117 | -106.0339
26.0154 | -106.0343
26.0191 | -106.0344
26.0227 | -106.0361
26.0264 | -106.0369
26.0301 | -106.0373
26.0337 | -106.0367
26.0374 | -106.0366
26.0411 | -106.0385
26.0447 | -106.042
26.0484 | -106.0452
26.0521 | -106.0492
26.0557 | -106.0513
26.0594 | -106.0533
26.0631 | -106.0549
26.0667 | -106.0568
26.0704 | -106.0581
26.074 | -106.0585
26.0777 | -106.0609
26.0814 | -106.0641
26.085 | -106.065
26.0887 | -106.0674
26.0924 | -106.0717
26.096 | -106.0743
26.0997 | -106.0769
26.1034 | -106.0812
26.107 | -106.0863
26.1107 | -106.0918
26.1144 | -106.095
26.118 | -106.0974
26.1217 | -106.1009
26.1254 | -106.1056
26.129 | -106.1101
26.1327 | -106.1143
26.1364 | -106.1173
26.14 | -106.1203
26.1437 | -106.1242
26.1474 | -106.1272
26.151 | -106.129
26.1547 | -106.1311
26.1584 | -106.1331
26.162 | -106.1355
26.1657 | -106.1375
26.1694 | -106.14
26.173 | -106.1401
26.1767 | -106.1396
26.1804 | -106.1393
26.184 | -106.1391
26.1877 | -106.139
26.1913 | -106.1378
26.195 | -106.1357
26.1987 | -106.133
26.2023 | -106.1307
26.206 | -106.129
26.2097 | -106.1265
26.2133 | -106.1238
26.217 | -106.1191
26.2207 | -106.1147
26.2243 | -106.109
26.228 | -106.1037
26.2317 | -106.0998
26.2353 | -106.0947
26.239 | -106.0893
26.2427 | -106.0828
26.2463 | -106.076
26.25 | -106.0695
26.2537 | -106.0628
26.2573 | -106.0567
26.261 | -106.05
26.2647 | -106.0428
26.2683 | -106.0343
26.272 | -106.0262
26.2757 | -106.0174
26.2793 | -106.0083
26.283 | -105.9998
26.2867 | -105.9905
26.2903 | -105.9818
26.294 | -105.9768
26.2977 | -105.9702
26.3013 | -105.9627
26.305 | -105.9559
26.3087 | -105.9488
26.3123 | -105.9421
26.316 | -105.9357
26.3196 | -105.9317
26.3233 | -105.9283
26.327 | -105.926
26.3306 | -105.9239
26.3343 | -105.9213
26.338 | -105.9191
26.3416 | -105.9183
26.3453 | -105.9199
26.349 | -105.9213
26.3526 | -105.9249
26.3563 | -105.9279
26.36 | -105.9316
26.3636 | -105.9343
26.3673 | -105.9366
26.371 | -105.9407
26.3746 | -105.9467
26.3783 | -105.9531
26.382 | -105.9596
26.3856 | -105.9645
26.3893 | -105.9694
26.393 | -105.9754
26.3966 | -105.9818
26.4003 | -105.9874
26.404 | -105.9951
26.4076 | -106.0032
26.4113 | -106.009
26.415 | -106.0145
26.4186 | -106.02
26.4223 | -106.026
26.426 | -106.0341
26.4296 | -106.041
26.4333 | -106.0474
26.4369 | -106.0536
26.4406 | -106.0588
26.4443 | -106.0634
26.4479 | -106.0689
26.4516 | -106.0736
26.4553 | -106.0774
26.4589 | -106.0833
26.4626 | -106.0881
26.4663 | -106.0912
26.4699 | -106.0947
26.4736 | -106.0987
26.4773 | -106.1003
26.4809 | -106.098
26.4846 | -106.096
26.4883 | -106.0964
26.4919 | -106.0974
26.4956 | -106.0972
26.4993 | -106.0963
26.5029 | -106.0963
26.5066 | -106.0945
26.5103 | -106.0925
26.5139 | -106.09
26.5176 | -106.0862
26.5213 | -106.0839
26.5249 | -106.0821
26.5286 | -106.079
26.5323 | -106.076
26.5359 | -106.0732
26.5396 | -106.069
26.5433 | -106.0645
26.5469 | -106.0614
26.5506 | -106.0571
26.5543 | -106.0515
26.5579 | -106.0445
26.5616 | -106.0356
26.5652 | -106.0276
26.5689 | -106.0202
26.5726 | -106.0137
26.5762 | -106.0074
26.5799 | -105.9997
26.5836 | -105.9926
26.5872 | -105.9848
26.5909 | -105.9789
26.5946 | -105.9736
26.5982 | -105.9679
26.6019 | -105.9628
26.6056 | -105.9581
26.6092 | -105.9535
26.6129 | -105.9489
26.6166 | -105.9443
26.6202 | -105.9403
26.6239 | -105.9365
26.6276 | -105.9323
26.6312 | -105.9291
26.6349 | -105.9262
26.6386 | -105.925
26.6422 | -105.9249
26.6459 | -105.9244
26.6496 | -105.9242
26.6532 | -105.9259
26.6569 | -105.928
26.6606 | -105.9312
26.6642 | -105.9358
26.6679 | -105.9373
26.6716 | -105.9397
26.6752 | -105.9426
26.6789 | -105.947
26.6826 | -105.9521
26.6862 | -105.9554
26.6899 | -105.9597
26.6935 | -105.9637
26.6972 | -105.9666
26.7009 | -105.9693
26.7045 | -105.9728
26.7082 | -105.9779
26.7119 | -105.9844
26.7155 | -105.9907
26.7192 | -105.9958
26.7229 | -106.0016
26.7265 | -106.0076
26.7302 | -106.013
26.7339 | -106.0178
26.7375 | -106.0243
26.7412 | -106.0316
26.7449 | -106.0386
26.7485 | -106.0442
26.7522 | -106.0495
26.7559 | -106.0555
26.7595 | -106.0622
26.7632 | -106.0686
26.7669 | -106.0736
26.7705 | -106.0787
26.7742 | -106.0835
26.7779 | -106.0891
26.7815 | -106.0949
26.7852 | -106.1017
26.7889 | -106.1071
26.7925 | -106.1124
26.7962 | -106.1166
26.7999 | -106.1189
26.8035 | -106.1212
26.8072 | -106.1246
26.8108 | -106.1281
26.8145 | -106.1302
26.8182 | -106.1329
26.8218 | -106.136
26.8255 | -106.1369
26.8292 | -106.1374
26.8328 | -106.1391
26.8365 | -106.142
26.8402 | -106.1435
26.8438 | -106.1447
26.8475 | -106.1469
26.8512 | -106.1487
26.8548 | -106.1522
26.8585 | -106.1573
26.8622 | -106.1603
26.8658 | -106.162
26.8695 | -106.162
26.8732 | -106.1625
26.8768 | -106.1625
26.8805 | -106.1626
26.8842 | -106.1642
26.8878 | -106.166
26.8915 | -106.1688
26.8952 | -106.1711
26.8988 | -106.1713
26.9025 | -106.1726
26.9062 | -106.1749
26.9098 | -106.1755
26.9135 | -106.175
26.9172 | -106.1757
26.9208 | -106.1775
26.9245 | -106.1788
26.9282 | -106.1814
26.9318 | -106.1842
26.9355 | -106.1875
26.9392 | -106.1921
26.9428 | -106.1957
26.9465 | -106.1992
26.9501 | -106.2019
26.9538 | -106.2051
26.9575 | -106.2087
26.9611 | -106.2119
26.9648 | -106.2147
26.9685 | -106.2128
26.9721 | -106.2138
26.9758 | -106.216
26.9795 | -106.2183
26.9831 | -106.2197
26.9868 | -106.2211
26.9905 | -106.2232
26.9941 | -106.2261
26.9978 | -106.2295
27.0015 | -106.2319
27.0051 | -106.2326
27.0088 | -106.2351
27.0125 | -106.2375
27.0161 | -106.2398
27.0198 | -106.2403
27.0235 | -106.2407
27.0271 | -106.2425
27.0308 | -106.2432
27.0345 | -106.2437
27.0381 | -106.2436
27.0418 | -106.2422
27.0455 | -106.2408
27.0491 | -106.2415
27.0528 | -106.2415
27.0565 | -106.2404
27.0601 | -106.2385
27.0638 | -106.2383
27.0674 | -106.2379
27.0711 | -106.2364
27.0748 | -106.2349
27.0784 | -106.2326
27.0821 | -106.2308
27.0858 | -106.2275
27.0894 | -106.2236
27.0931 | -106.2204
27.0968 | -106.2162
27.1004 | -106.2119
27.1041 | -106.2079
27.1078 | -106.2026
27.1114 | -106.1975
27.1151 | -106.1929
27.1188 | -106.1889
27.1224 | -106.1843
27.1261 | -106.1783
27.1298 | -106.1714
27.1334 | -106.1657
27.1371 | -106.1606
27.1408 | -106.1561
27.1444 | -106.1528
27.1481 | -106.1485
27.1518 | -106.1438
27.1554 | -106.138
27.1591 | -106.1297
27.1628 | -106.1211
27.1664 | -106.1111
27.1701 | -106.1019
27.1738 | -106.0939
27.1774 | -106.0868
27.1811 | -106.0799
27.1848 | -106.0742
27.1884 | -106.0679
27.1921 | -106.0609
27.1957 | -106.0533
27.1994 | -106.0459
27.2031 | -106.0407
27.2067 | -106.0341
27.2104 | -106.0266
27.2141 | -106.0192
27.2177 | -106.0112
27.2214 | -106.004
27.2251 | -105.996
27.2287 | -105.9873
27.2324 | -105.9792
27.2361 | -105.9724
27.2397 | -105.9672
27.2434 | -105.9611
27.2471 | -105.9548
27.2507 | -105.95
27.2544 | -105.9438
27.2581 | -105.9372
27.2617 | -105.9307
27.2654 | -105.9246
27.2691 | -105.9183
27.2727 | -105.912
27.2764 | -105.9059
27.2801 | -105.9006
27.2837 | -105.8948
27.2874 | -105.8886
27.2911 | -105.8837
27.2947 | -105.8797
27.2984 | -105.8765
27.3021 | -105.8728
27.3057 | -105.8697
27.3094 | -105.8656
27.3131 | -105.8623
27.3167 | -105.859
27.3204 | -105.8546
27.324 | -105.8531
27.3277 | -105.8525
27.3314 | -105.8518
27.335 | -105.8509
27.3387 | -105.8499
27.3424 | -105.8491
27.346 | -105.8492
27.3497 | -105.85
27.3534 | -105.8506
27.357 | -105.8524
27.3607 | -105.8514
27.3644 | -105.8507
27.368 | -105.852
27.3717 | -105.8544
27.3754 | -105.8561
27.379 | -105.8583
27.3827 | -105.859
27.3864 | -105.8593
27.39 | -105.8622
27.3937 | -105.8654
27.3974 | -105.8682
27.401 | -105.8705
27.4047 | -105.8734
27.4084 | -105.8775
27.412 | -105.881
27.4157 | -105.8845
27.4194 | -105.8853
27.423 | -105.8867
27.4267 | -105.8905
27.4304 | -105.8947
27.434 | -105.8967
27.4377 | -105.8986
27.4413 | -105.9027
27.445 | -105.9062
27.4487 | -105.9076
27.4523 | -105.9107
27.456 | -105.9131
27.4597 | -105.9129
27.4633 | -105.914
27.467 | -105.9149
27.4707 | -105.9165
27.4743 | -105.9173
27.478 | -105.9176
27.4817 | -105.9181
27.4853 | -105.9181
27.489 | -105.9167
27.4927 | -105.9155
27.4963 | -105.9085
27.5 | -105.9071
27.5037 | -105.9058
27.5073 | -105.9038
27.511 | -105.9017
27.5147 | -105.9002
27.5183 | -105.8991
27.522 | -105.897
27.5257 | -105.8937
27.5293 | -105.8897
27.533 | -105.8851
27.5367 | -105.881
27.5403 | -105.8776
27.544 | -105.8756
27.5477 | -105.8739
27.5513 | -105.8735
27.555 | -105.8745
27.5587 | -105.8732
27.5623 | -105.8682
27.566 | -105.8634
27.5696 | -105.8602
27.5733 | -105.8579
27.577 | -105.856
27.5806 | -105.8529
27.5843 | -105.85
27.588 | -105.8466
27.5916 | -105.8434
27.5953 | -105.8424
27.599 | -105.8409
27.6026 | -105.8396
27.6063 | -105.838
27.61 | -105.836
27.6136 | -105.8343
27.6173 | -105.8324
27.621 | -105.8294
27.6246 | -105.8275
27.6283 | -105.8266
27.632 | -105.8267
27.6356 | -105.8258
27.6393 | -105.8244
27.643 | -105.8244
27.6466 | -105.8244
27.6503 | -105.8244
27.654 | -105.8236
27.6576 | -105.8233
27.6613 | -105.8226
27.665 | -105.8226
27.6686 | -105.8234
27.6723 | -105.8229
27.676 | -105.8229
27.6796 | -105.8216
27.6833 | -105.8198
27.6869 | -105.8187
27.6906 | -105.8191
27.6943 | -105.8198
27.6979 | -105.82
27.7016 | -105.8189
27.7053 | -105.8189
27.7089 | -105.8204
27.7126 | -105.8185
27.7163 | -105.817
27.7199 | -105.8163
27.7236 | -105.8166
27.7273 | -105.817
27.7309 | -105.8175
27.7346 | -105.8169
27.7383 | -105.8177
27.7419 | -105.8194
27.7456 | -105.8185
27.7493 | -105.8185
27.7529 | -105.8188
27.7566 | -105.8175
27.7603 | -105.8167
27.7639 | -105.8161
27.7676 | -105.816
27.7713 | -105.8162
27.7749 | -105.8171
27.7786 | -105.8155
27.7823 | -105.8151
27.7859 | -105.8138
27.7896 | -105.8128
27.7933 | -105.812
27.7969 | -105.8105
27.8006 | -105.8081
27.8043 | -105.8065
27.8079 | -105.8068
27.8116 | -105.8051
27.8152 | -105.8019
27.8189 | -105.799
27.8226 | -105.7966
27.8262 | -105.7944
27.8299 | -105.7905
27.8336 | -105.7873
27.8372 | -105.7846
27.8409 | -105.7814
27.8446 | -105.7775
27.8482 | -105.7748
27.8519 | -105.7724
27.8556 | -105.7702
27.8592 | -105.7676
27.8629 | -105.7642
27.8666 | -105.7601
27.8702 | -105.7565
27.8739 | -105.7543
27.8776 | -105.7513
27.8812 | -105.7485
27.8849 | -105.7474
27.8886 | -105.7462
27.8922 | -105.744
27.8959 | -105.7413
27.8996 | -105.7384
27.9032 | -105.7367
27.9069 | -105.7353
27.9106 | -105.7349
27.9142 | -105.7346
27.9179 | -105.7342
27.9216 | -105.7345
27.9252 | -105.7358
27.9289 | -105.7384
27.9326 | -105.7399
27.9362 | -105.7407
27.9399 | -105.7427
27.9435 | -105.7445
27.9472 | -105.7455
27.9509 | -105.7481
27.9545 | -105.7532
27.9582 | -105.7573
27.9619 | -105.7602
27.9655 | -105.7614
27.9692 | -105.764
27.9729 | -105.7683
27.9765 | -105.7737
27.9802 | -105.7788
27.9839 | -105.7827
27.9875 | -105.7868
27.9912 | -105.7931
27.9949 | -105.7976
27.9985 | -105.8017
28.0022 | -105.8073