## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 213 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 211
Hood | 02-92B/P | 150
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1006.9 | hPa |  | PRC PRO 01
Humidity | 37.8 | %/Hr
Internal temperature of calibration bench | 24.1 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1019 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.3 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.96 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.02 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3136 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 05/04/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 12/04/2018
Engineer | Anthony HUE | Approval |  | 23/04/2018
CEO | Robert GLEMOT | Approval |  | 23/04/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3015 | 57.4687
0.012209 | -35.874 | 38.2981
0.018313 | -34.8113 | 28.1611
0.024417 | -34.3521 | 21.9228
0.030521 | -34.2393 | 18.1682
0.036626 | -34.0952 | 15.119
0.04273 | -33.9977 | 12.7521
0.048834 | -33.9523 | 11.2461
0.054938 | -33.8883 | 10.1212
0.061043 | -33.8292 | 9.0415
0.067147 | -33.8192 | 8.2088
0.073251 | -33.8194 | 7.4433
0.079355 | -33.8366 | 6.7757
0.08546 | -33.8497 | 6.1742
0.091564 | -33.8448 | 5.7342
0.097668 | -33.8277 | 5.4927
0.10377 | -33.8079 | 5.2479
0.10988 | -33.7924 | 4.9737
0.11598 | -33.7839 | 4.6825
0.12209 | -33.7852 | 4.3972
0.12819 | -33.7937 | 4.109
0.13429 | -33.8027 | 3.8179
0.1404 | -33.8055 | 3.547
0.1465 | -33.8071 | 3.2946
0.15261 | -33.8047 | 3.0786
0.15871 | -33.7989 | 2.9169
0.16482 | -33.7914 | 2.8111
0.17092 | -33.7897 | 2.7689
0.17702 | -33.7881 | 2.6955
0.18313 | -33.7874 | 2.6177
0.18923 | -33.7869 | 2.5098
0.19534 | -33.785 | 2.3836
0.20144 | -33.783 | 2.241
0.20754 | -33.7798 | 2.1172
0.21365 | -33.7786 | 2.0011
0.21975 | -33.7781 | 1.9026
0.22586 | -33.7792 | 1.8156
0.23196 | -33.7818 | 1.7455
0.23807 | -33.7858 | 1.6857
0.24417 | -33.7896 | 1.6288
0.25027 | -33.7933 | 1.5676
0.25638 | -33.7953 | 1.506
0.26248 | -33.7966 | 1.439
0.26859 | -33.7969 | 1.3725
0.27469 | -33.7967 | 1.3005
0.2808 | -33.7974 | 1.227
0.2869 | -33.7979 | 1.1559
0.293 | -33.7987 | 1.0887
0.29911 | -33.7996 | 1.0276
0.30521 | -33.8006 | 0.96716
0.31132 | -33.8019 | 0.92618
0.31742 | -33.8035 | 0.8814
0.32353 | -33.8043 | 0.84466
0.32963 | -33.8057 | 0.81413
0.33573 | -33.8073 | 0.79223
0.34184 | -33.8093 | 0.76975
0.34794 | -33.8121 | 0.7447
0.35405 | -33.8148 | 0.71936
0.36015 | -33.8178 | 0.69271
0.36626 | -33.8203 | 0.66381
0.37236 | -33.8229 | 0.62558
0.37846 | -33.8254 | 0.57915
0.38457 | -33.8263 | 0.5194
0.39067 | -33.8274 | 0.47062
0.39678 | -33.8281 | 0.41698
0.40288 | -33.8284 | 0.36125
0.40899 | -33.8281 | 0.30826
0.41509 | -33.8272 | 0.2574
0.42119 | -33.8268 | 0.21491
0.4273 | -33.8266 | 0.16884
0.4334 | -33.8262 | 0.12284
0.43951 | -33.8262 | 0.080238
0.44561 | -33.8269 | 0.041336
0.45172 | -33.8285 | 0.010027
0.45782 | -33.8292 | -0.026391
0.46392 | -33.8305 | -0.06231
0.47003 | -33.8316 | -0.093956
0.47613 | -33.8326 | -0.12769
0.48224 | -33.8332 | -0.1618
0.48834 | -33.8324 | -0.19552
0.49445 | -33.8324 | -0.22904
0.50055 | -33.8319 | -0.2602
0.50665 | -33.8313 | -0.29295
0.51276 | -33.8303 | -0.32571
0.51886 | -33.8293 | -0.35801
0.52497 | -33.828 | -0.38027
0.53107 | -33.8271 | -0.40862
0.53717 | -33.8264 | -0.43605
0.54328 | -33.8253 | -0.45804
0.54938 | -33.8248 | -0.47979
0.55549 | -33.8246 | -0.49982
0.56159 | -33.8254 | -0.51099
0.5677 | -33.8255 | -0.52545
0.5738 | -33.8263 | -0.54129
0.5799 | -33.8267 | -0.55613
0.58601 | -33.8276 | -0.56609
0.59211 | -33.8291 | -0.57217
0.59822 | -33.8303 | -0.58194
0.60432 | -33.8316 | -0.59415
0.61043 | -33.8327 | -0.60143
0.61653 | -33.8338 | -0.61121
0.62263 | -33.8353 | -0.61997
0.62874 | -33.8361 | -0.63121
0.63484 | -33.837 | -0.64002
0.64095 | -33.8376 | -0.65056
0.64705 | -33.8381 | -0.66091
0.65316 | -33.8383 | -0.67361
0.65926 | -33.8386 | -0.68956
0.66536 | -33.8377 | -0.70773
0.67147 | -33.8376 | -0.72174
0.67757 | -33.8373 | -0.73659
0.68368 | -33.8368 | -0.7566
0.68978 | -33.8362 | -0.77432
0.69589 | -33.8355 | -0.79394
0.70199 | -33.8347 | -0.81742
0.70809 | -33.8341 | -0.83727
0.7142 | -33.8336 | -0.85823
0.7203 | -33.8327 | -0.87796
0.72641 | -33.8321 | -0.90098
0.73251 | -33.8316 | -0.92133
0.73862 | -33.8319 | -0.94353
0.74472 | -33.8316 | -0.9642
0.75082 | -33.8314 | -0.9873
0.75693 | -33.8318 | -1.0102
0.76303 | -33.8323 | -1.0328
0.76914 | -33.8342 | -1.0537
0.77524 | -33.8352 | -1.0741
0.78135 | -33.8365 | -1.0943
0.78745 | -33.8379 | -1.1151
0.79355 | -33.8391 | -1.1354
0.79966 | -33.8407 | -1.1557
0.80576 | -33.8428 | -1.1765
0.81187 | -33.8443 | -1.1988
0.81797 | -33.8462 | -1.2207
0.82408 | -33.8478 | -1.242
0.83018 | -33.8494 | -1.2637
0.83628 | -33.8512 | -1.2833
0.84239 | -33.8523 | -1.3057
0.84849 | -33.8536 | -1.3284
0.8546 | -33.8549 | -1.3494
0.8607 | -33.856 | -1.3716
0.86681 | -33.857 | -1.3918
0.87291 | -33.8579 | -1.412
0.87901 | -33.8583 | -1.4316
0.88512 | -33.859 | -1.4514
0.89122 | -33.8598 | -1.4711
0.89733 | -33.8601 | -1.4879
0.90343 | -33.8604 | -1.5063
0.90953 | -33.8606 | -1.5207
0.91564 | -33.8606 | -1.5372
0.92174 | -33.8608 | -1.553
0.92785 | -33.8609 | -1.5683
0.93395 | -33.8611 | -1.5823
0.94006 | -33.8614 | -1.5957
0.94616 | -33.8613 | -1.6071
0.95226 | -33.8615 | -1.6202
0.95837 | -33.8618 | -1.6335
0.96447 | -33.8622 | -1.646
0.97058 | -33.8621 | -1.6586
0.97668 | -33.8621 | -1.6712
0.98279 | -33.8621 | -1.683
0.98889 | -33.8622 | -1.6926
0.99499 | -33.8624 | -1.7049
1.0011 | -33.8624 | -1.7145
1.0072 | -33.8623 | -1.7249
1.0133 | -33.8623 | -1.7341
1.0194 | -33.8622 | -1.7459
1.0255 | -33.862 | -1.7582
1.0316 | -33.8618 | -1.7693
1.0377 | -33.8617 | -1.7823
1.0438 | -33.8613 | -1.7939
1.0499 | -33.8608 | -1.8082
1.056 | -33.8604 | -1.8219
1.0621 | -33.86 | -1.8354
1.0682 | -33.8597 | -1.847
1.0743 | -33.8593 | -1.8619
1.0805 | -33.8589 | -1.8757
1.0866 | -33.8581 | -1.8917
1.0927 | -33.8574 | -1.9058
1.0988 | -33.8569 | -1.9205
1.1049 | -33.8562 | -1.9353
1.111 | -33.8555 | -1.9511
1.1171 | -33.855 | -1.966
1.1232 | -33.8544 | -1.9821
1.1293 | -33.8541 | -1.9969
1.1354 | -33.8539 | -2.0126
1.1415 | -33.8536 | -2.029
1.1476 | -33.853 | -2.0456
1.1537 | -33.8525 | -2.0615
1.1598 | -33.8525 | -2.0777
1.1659 | -33.8523 | -2.0919
1.172 | -33.852 | -2.1072
1.1781 | -33.8517 | -2.1228
1.1842 | -33.8515 | -2.1379
1.1903 | -33.8517 | -2.1505
1.1964 | -33.8515 | -2.1655
1.2025 | -33.8515 | -2.181
1.2086 | -33.8517 | -2.1949
1.2147 | -33.8516 | -2.2092
1.2209 | -33.8519 | -2.2218
1.227 | -33.8525 | -2.2342
1.2331 | -33.8528 | -2.2475
1.2392 | -33.853 | -2.2608
1.2453 | -33.8533 | -2.2734
1.2514 | -33.8537 | -2.2849
1.2575 | -33.8541 | -2.2975
1.2636 | -33.8549 | -2.3074
1.2697 | -33.8554 | -2.3176
1.2758 | -33.856 | -2.3297
1.2819 | -33.8567 | -2.3408
1.288 | -33.8573 | -2.3531
1.2941 | -33.8582 | -2.3628
1.3002 | -33.8589 | -2.372
1.3063 | -33.8594 | -2.3821
1.3124 | -33.8601 | -2.3924
1.3185 | -33.8608 | -2.4027
1.3246 | -33.8614 | -2.4127
1.3307 | -33.8623 | -2.4219
1.3368 | -33.8632 | -2.4312
1.3429 | -33.8638 | -2.4409
1.349 | -33.8643 | -2.4507
1.3551 | -33.865 | -2.4607
1.3613 | -33.8656 | -2.4694
1.3674 | -33.8662 | -2.4796
1.3735 | -33.8667 | -2.49
1.3796 | -33.8672 | -2.4996
1.3857 | -33.8676 | -2.5095
1.3918 | -33.8682 | -2.5203
1.3979 | -33.8688 | -2.5299
1.404 | -33.8692 | -2.5413
1.4101 | -33.8697 | -2.5521
1.4162 | -33.87 | -2.5628
1.4223 | -33.8705 | -2.5746
1.4284 | -33.871 | -2.5849
1.4345 | -33.8712 | -2.5967
1.4406 | -33.8715 | -2.6099
1.4467 | -33.8719 | -2.6217
1.4528 | -33.8724 | -2.6331
1.4589 | -33.8728 | -2.6448
1.465 | -33.873 | -2.6555
1.4711 | -33.8735 | -2.6669
1.4772 | -33.874 | -2.6799
1.4833 | -33.8744 | -2.6929
1.4894 | -33.8749 | -2.7049
1.4955 | -33.8754 | -2.7165
1.5016 | -33.8759 | -2.729
1.5078 | -33.8762 | -2.7423
1.5139 | -33.8766 | -2.7551
1.52 | -33.877 | -2.7672
1.5261 | -33.8773 | -2.7794
1.5322 | -33.8777 | -2.7931
1.5383 | -33.878 | -2.8051
1.5444 | -33.8785 | -2.8172
1.5505 | -33.879 | -2.8302
1.5566 | -33.8795 | -2.8426
1.5627 | -33.8799 | -2.8557
1.5688 | -33.8805 | -2.8688
1.5749 | -33.8809 | -2.881
1.581 | -33.8812 | -2.8936
1.5871 | -33.8815 | -2.9067
1.5932 | -33.8819 | -2.9187
1.5993 | -33.8822 | -2.9322
1.6054 | -33.8826 | -2.945
1.6115 | -33.8831 | -2.9587
1.6176 | -33.8834 | -2.9708
1.6237 | -33.8836 | -2.9839
1.6298 | -33.8839 | -2.9957
1.6359 | -33.8843 | -3.0079
1.642 | -33.8847 | -3.0191
1.6482 | -33.885 | -3.0316
1.6543 | -33.8854 | -3.0447
1.6604 | -33.8858 | -3.0561
1.6665 | -33.8859 | -3.068
1.6726 | -33.8861 | -3.0801
1.6787 | -33.8865 | -3.0921
1.6848 | -33.8865 | -3.1041
1.6909 | -33.8869 | -3.1161
1.697 | -33.8872 | -3.1279
1.7031 | -33.8873 | -3.1398
1.7092 | -33.8874 | -3.1515
1.7153 | -33.8876 | -3.1628
1.7214 | -33.8876 | -3.1746
1.7275 | -33.8876 | -3.1865
1.7336 | -33.8876 | -3.1991
1.7397 | -33.8879 | -3.211
1.7458 | -33.8879 | -3.2232
1.7519 | -33.8881 | -3.2349
1.758 | -33.8882 | -3.2467
1.7641 | -33.8884 | -3.258
1.7702 | -33.8884 | -3.2709
1.7763 | -33.8885 | -3.2821
1.7824 | -33.8886 | -3.2944
1.7885 | -33.8886 | -3.3069
1.7947 | -33.8888 | -3.3188
1.8008 | -33.8889 | -3.3298
1.8069 | -33.889 | -3.3412
1.813 | -33.8891 | -3.3535
1.8191 | -33.8891 | -3.3664
1.8252 | -33.8893 | -3.3793
1.8313 | -33.8894 | -3.3917
1.8374 | -33.8896 | -3.4036
1.8435 | -33.8898 | -3.4154
1.8496 | -33.8898 | -3.4284
1.8557 | -33.89 | -3.4409
1.8618 | -33.8903 | -3.4536
1.8679 | -33.8905 | -3.4656
1.874 | -33.8906 | -3.4779
1.8801 | -33.8908 | -3.4904
1.8862 | -33.8911 | -3.5025
1.8923 | -33.8912 | -3.5149
1.8984 | -33.8913 | -3.5274
1.9045 | -33.8914 | -3.5392
1.9106 | -33.8914 | -3.5517
1.9167 | -33.8915 | -3.5642
1.9228 | -33.8917 | -3.5765
1.9289 | -33.8917 | -3.5888
1.9351 | -33.8919 | -3.6001
1.9412 | -33.8921 | -3.6116
1.9473 | -33.8925 | -3.6241
1.9534 | -33.8926 | -3.6359
1.9595 | -33.8928 | -3.6481
1.9656 | -33.8931 | -3.6609
1.9717 | -33.8933 | -3.6728
1.9778 | -33.8936 | -3.6847
1.9839 | -33.8938 | -3.6964
1.99 | -33.894 | -3.7078
1.9961 | -33.894 | -3.7204
2.0022 | -33.8942 | -3.7323
2.0083 | -33.8943 | -3.7439
2.0144 | -33.8945 | -3.7558
2.0205 | -33.8949 | -3.7679
2.0266 | -33.8952 | -3.7797
2.0327 | -33.8954 | -3.7916
2.0388 | -33.8958 | -3.8037
2.0449 | -33.896 | -3.8157
2.051 | -33.8963 | -3.8276
2.0571 | -33.8966 | -3.8395
2.0632 | -33.8967 | -3.8513
2.0693 | -33.8969 | -3.8641
2.0754 | -33.8972 | -3.8758
2.0816 | -33.8975 | -3.8883
2.0877 | -33.8976 | -3.9003
2.0938 | -33.898 | -3.9129
2.0999 | -33.8982 | -3.925
2.106 | -33.8984 | -3.937
2.1121 | -33.8986 | -3.9494
2.1182 | -33.8988 | -3.9605
2.1243 | -33.8991 | -3.9715
2.1304 | -33.8993 | -3.9831
2.1365 | -33.8996 | -3.9953
2.1426 | -33.8999 | -4.0077
2.1487 | -33.9001 | -4.0193
2.1548 | -33.9006 | -4.0316
2.1609 | -33.901 | -4.0434
2.167 | -33.9014 | -4.055
2.1731 | -33.9017 | -4.0675
2.1792 | -33.9019 | -4.0789
2.1853 | -33.9023 | -4.0914
2.1914 | -33.9026 | -4.1031
2.1975 | -33.903 | -4.1151
2.2036 | -33.9032 | -4.1275
2.2097 | -33.9035 | -4.1393
2.2158 | -33.9039 | -4.1503
2.222 | -33.9042 | -4.1612
2.2281 | -33.9044 | -4.173
2.2342 | -33.9048 | -4.1847
2.2403 | -33.905 | -4.1958
2.2464 | -33.9053 | -4.2069
2.2525 | -33.9056 | -4.2186
2.2586 | -33.9059 | -4.2296
2.2647 | -33.9062 | -4.2403
2.2708 | -33.9065 | -4.2519
2.2769 | -33.9069 | -4.263
2.283 | -33.9071 | -4.2745
2.2891 | -33.9073 | -4.2859
2.2952 | -33.9075 | -4.2967
2.3013 | -33.9077 | -4.3079
2.3074 | -33.908 | -4.3186
2.3135 | -33.9082 | -4.3303
2.3196 | -33.9084 | -4.3408
2.3257 | -33.9086 | -4.3517
2.3318 | -33.9087 | -4.3633
2.3379 | -33.9089 | -4.3744
2.344 | -33.9091 | -4.3861
2.3501 | -33.9093 | -4.397
2.3562 | -33.9094 | -4.4076
2.3623 | -33.9095 | -4.4183
2.3685 | -33.9096 | -4.4292
2.3746 | -33.9098 | -4.4408
2.3807 | -33.91 | -4.4516
2.3868 | -33.9101 | -4.4625
2.3929 | -33.9103 | -4.4734
2.399 | -33.9103 | -4.4842
2.4051 | -33.9106 | -4.495
2.4112 | -33.9108 | -4.5059
2.4173 | -33.9109 | -4.5169
2.4234 | -33.9109 | -4.528
2.4295 | -33.9111 | -4.5392
2.4356 | -33.9112 | -4.5505
2.4417 | -33.9113 | -4.5618
2.4478 | -33.9113 | -4.573
2.4539 | -33.9115 | -4.5837
2.46 | -33.9116 | -4.5955
2.4661 | -33.9117 | -4.6068
2.4722 | -33.9118 | -4.6173
2.4783 | -33.9118 | -4.628
2.4844 | -33.9119 | -4.6393
2.4905 | -33.912 | -4.6505
2.4966 | -33.9121 | -4.6616
2.5027 | -33.9121 | -4.6723
2.5089 | -33.9122 | -4.6835
2.515 | -33.9122 | -4.6946
2.5211 | -33.9123 | -4.7057
2.5272 | -33.9123 | -4.7168
2.5333 | -33.9124 | -4.728
2.5394 | -33.9125 | -4.739
2.5455 | -33.9126 | -4.7507
2.5516 | -33.9127 | -4.7612
2.5577 | -33.9129 | -4.773
2.5638 | -33.913 | -4.7849
2.5699 | -33.913 | -4.7965
2.576 | -33.9131 | -4.8083
2.5821 | -33.9131 | -4.82
2.5882 | -33.9132 | -4.8312
2.5943 | -33.9133 | -4.8422
2.6004 | -33.9135 | -4.8534
2.6065 | -33.9136 | -4.864
2.6126 | -33.9137 | -4.8746
2.6187 | -33.9138 | -4.8854
2.6248 | -33.9139 | -4.8964
2.6309 | -33.9141 | -4.9081
2.637 | -33.9144 | -4.9199
2.6431 | -33.9145 | -4.9306
2.6492 | -33.9146 | -4.9422
2.6554 | -33.9148 | -4.9532
2.6615 | -33.9149 | -4.9642
2.6676 | -33.9151 | -4.9758
2.6737 | -33.9154 | -4.9871
2.6798 | -33.9157 | -4.9984
2.6859 | -33.9159 | -5.0098
2.692 | -33.916 | -5.0208
2.6981 | -33.9163 | -5.032
2.7042 | -33.9165 | -5.0432
2.7103 | -33.9167 | -5.0541
2.7164 | -33.917 | -5.0652
2.7225 | -33.9173 | -5.0766
2.7286 | -33.9175 | -5.0878
2.7347 | -33.9177 | -5.0985
2.7408 | -33.9181 | -5.1087
2.7469 | -33.9183 | -5.12
2.753 | -33.9186 | -5.1306
2.7591 | -33.9189 | -5.1421
2.7652 | -33.9192 | -5.1536
2.7713 | -33.9193 | -5.1645
2.7774 | -33.9197 | -5.1755
2.7835 | -33.9198 | -5.1865
2.7896 | -33.9201 | -5.1971
2.7958 | -33.9204 | -5.208
2.8019 | -33.9207 | -5.2183
2.808 | -33.921 | -5.2292
2.8141 | -33.9214 | -5.2404
2.8202 | -33.9217 | -5.2513
2.8263 | -33.9221 | -5.2623
2.8324 | -33.9225 | -5.2737
2.8385 | -33.9229 | -5.2847
2.8446 | -33.9233 | -5.2953
2.8507 | -33.9237 | -5.3063
2.8568 | -33.9241 | -5.3168
2.8629 | -33.9244 | -5.3275
2.869 | -33.9248 | -5.3387
2.8751 | -33.9252 | -5.3495
2.8812 | -33.9256 | -5.3608
2.8873 | -33.926 | -5.3717
2.8934 | -33.9264 | -5.3828
2.8995 | -33.9267 | -5.394
2.9056 | -33.9271 | -5.4052
2.9117 | -33.9274 | -5.4163
2.9178 | -33.9279 | -5.4271
2.9239 | -33.9283 | -5.4386
2.93 | -33.9288 | -5.4501
2.9361 | -33.9291 | -5.461
2.9423 | -33.9295 | -5.4718
2.9484 | -33.93 | -5.4829
2.9545 | -33.9304 | -5.494
2.9606 | -33.9309 | -5.5055
2.9667 | -33.9313 | -5.5167
2.9728 | -33.9317 | -5.5281
2.9789 | -33.9322 | -5.5392
2.985 | -33.9327 | -5.5507
2.9911 | -33.9331 | -5.5621
2.9972 | -33.9336 | -5.5733
3.0033 | -33.934 | -5.5849
3.0094 | -33.9345 | -5.5961
3.0155 | -33.9349 | -5.6072
3.0216 | -33.9353 | -5.6186
3.0277 | -33.9357 | -5.6304
3.0338 | -33.9362 | -5.642
3.0399 | -33.9366 | -5.6534
3.046 | -33.937 | -5.6652
3.0521 | -33.9374 | -5.6764
3.0582 | -33.9377 | -5.6886
3.0643 | -33.9381 | -5.7003
3.0704 | -33.9385 | -5.7123
3.0765 | -33.9388 | -5.7235
3.0827 | -33.9391 | -5.7349
3.0888 | -33.9395 | -5.7463
3.0949 | -33.94 | -5.7579
3.101 | -33.9403 | -5.7696
3.1071 | -33.9407 | -5.7815
3.1132 | -33.9411 | -5.7935
3.1193 | -33.9415 | -5.8051
3.1254 | -33.942 | -5.8167
3.1315 | -33.9424 | -5.8289
3.1376 | -33.9427 | -5.8405
3.1437 | -33.9431 | -5.8522
3.1498 | -33.9434 | -5.8645
3.1559 | -33.9438 | -5.8762
3.162 | -33.9441 | -5.8885
3.1681 | -33.9444 | -5.9007
3.1742 | -33.9447 | -5.9125
3.1803 | -33.9451 | -5.9243
3.1864 | -33.9455 | -5.9359
3.1925 | -33.9458 | -5.9476
3.1986 | -33.9461 | -5.9598
3.2047 | -33.9464 | -5.9721
3.2108 | -33.9467 | -5.9834
3.2169 | -33.9471 | -5.9954
3.223 | -33.9473 | -6.0074
3.2292 | -33.9476 | -6.0194
3.2353 | -33.9478 | -6.0305
3.2414 | -33.9481 | -6.0422
3.2475 | -33.9484 | -6.0538
3.2536 | -33.9487 | -6.0654
3.2597 | -33.949 | -6.0766
3.2658 | -33.9492 | -6.0877
3.2719 | -33.9494 | -6.099
3.278 | -33.9497 | -6.1108
3.2841 | -33.95 | -6.1222
3.2902 | -33.9502 | -6.1338
3.2963 | -33.9505 | -6.1456
3.3024 | -33.9506 | -6.1569
3.3085 | -33.9509 | -6.1687
3.3146 | -33.9511 | -6.1801
3.3207 | -33.9514 | -6.1918
3.3268 | -33.9516 | -6.2035
3.3329 | -33.9518 | -6.2145
3.339 | -33.952 | -6.2255
3.3451 | -33.9522 | -6.2371
3.3512 | -33.9524 | -6.248
3.3573 | -33.9526 | -6.2596
3.3634 | -33.9528 | -6.2709
3.3696 | -33.953 | -6.2825
3.3757 | -33.9532 | -6.2935
3.3818 | -33.9534 | -6.3046
3.3879 | -33.9536 | -6.3159
3.394 | -33.9537 | -6.3274
3.4001 | -33.9539 | -6.3385
3.4062 | -33.954 | -6.3496
3.4123 | -33.9542 | -6.3605
3.4184 | -33.9544 | -6.372
3.4245 | -33.9546 | -6.383
3.4306 | -33.9547 | -6.3939
3.4367 | -33.9549 | -6.4046
3.4428 | -33.9551 | -6.4152
3.4489 | -33.9552 | -6.4264
3.455 | -33.9553 | -6.4374
3.4611 | -33.9556 | -6.4489
3.4672 | -33.9557 | -6.46
3.4733 | -33.9559 | -6.4713
3.4794 | -33.956 | -6.4823
3.4855 | -33.9562 | -6.4928
3.4916 | -33.9564 | -6.5034
3.4977 | -33.9565 | -6.5148
3.5038 | -33.9567 | -6.5259
3.5099 | -33.9568 | -6.5364
3.5161 | -33.9569 | -6.5469
3.5222 | -33.9571 | -6.5575
3.5283 | -33.9572 | -6.5684
3.5344 | -33.9573 | -6.5794
3.5405 | -33.9575 | -6.5901
3.5466 | -33.9576 | -6.6012
3.5527 | -33.9577 | -6.6121
3.5588 | -33.9578 | -6.6229
3.5649 | -33.9579 | -6.6337
3.571 | -33.958 | -6.6446
3.5771 | -33.9581 | -6.6559
3.5832 | -33.9583 | -6.6659
3.5893 | -33.9585 | -6.6767
3.5954 | -33.9587 | -6.6877
3.6015 | -33.9588 | -6.6988
3.6076 | -33.959 | -6.7096
3.6137 | -33.9592 | -6.7199
3.6198 | -33.9594 | -6.7303
3.6259 | -33.9596 | -6.7413
3.632 | -33.9597 | -6.7522
3.6381 | -33.9599 | -6.7631
3.6442 | -33.9601 | -6.7739
3.6503 | -33.9603 | -6.7842
3.6565 | -33.9605 | -6.7951
3.6626 | -33.9606 | -6.8059
3.6687 | -33.9608 | -6.8168
3.6748 | -33.9611 | -6.8275
3.6809 | -33.9613 | -6.8382
3.687 | -33.9614 | -6.8494
3.6931 | -33.9617 | -6.8598
3.6992 | -33.9619 | -6.8706
3.7053 | -33.9621 | -6.8813
3.7114 | -33.9622 | -6.8921
3.7175 | -33.9625 | -6.903
3.7236 | -33.9627 | -6.9132
3.7297 | -33.9629 | -6.9242
3.7358 | -33.9631 | -6.9352
3.7419 | -33.9633 | -6.9462
3.748 | -33.9636 | -6.9569
3.7541 | -33.9638 | -6.9677
3.7602 | -33.9641 | -6.9784
3.7663 | -33.9643 | -6.9889
3.7724 | -33.9645 | -6.9997
3.7785 | -33.9648 | -7.0103
3.7846 | -33.965 | -7.0209
3.7907 | -33.9653 | -7.0321
3.7969 | -33.9656 | -7.043
3.803 | -33.9658 | -7.0534
3.8091 | -33.966 | -7.0642
3.8152 | -33.9663 | -7.0748
3.8213 | -33.9666 | -7.0858
3.8274 | -33.9668 | -7.0963
3.8335 | -33.9672 | -7.1072
3.8396 | -33.9675 | -7.1181
3.8457 | -33.9678 | -7.129
3.8518 | -33.9681 | -7.1397
3.8579 | -33.9685 | -7.1505
3.864 | -33.9688 | -7.1615
3.8701 | -33.9691 | -7.1726
3.8762 | -33.9694 | -7.1833
3.8823 | -33.9696 | -7.1942
3.8884 | -33.97 | -7.2051
3.8945 | -33.9703 | -7.2162
3.9006 | -33.9706 | -7.2274
3.9067 | -33.9709 | -7.2386
3.9128 | -33.9712 | -7.2495
3.9189 | -33.9716 | -7.2602
3.925 | -33.9718 | -7.2712
3.9311 | -33.9721 | -7.2816
3.9372 | -33.9724 | -7.2927
3.9434 | -33.9727 | -7.3036
3.9495 | -33.9731 | -7.3146
3.9556 | -33.9735 | -7.3252
3.9617 | -33.9738 | -7.3357
3.9678 | -33.9741 | -7.3463
3.9739 | -33.9744 | -7.3572
3.98 | -33.9747 | -7.3682
3.9861 | -33.975 | -7.379
3.9922 | -33.9752 | -7.3899
3.9983 | -33.9756 | -7.4004
4.0044 | -33.9759 | -7.4108
4.0105 | -33.9762 | -7.4217
4.0166 | -33.9765 | -7.4326
4.0227 | -33.9769 | -7.4434
4.0288 | -33.9771 | -7.4547
4.0349 | -33.9774 | -7.4653
4.041 | -33.9777 | -7.4761
4.0471 | -33.9779 | -7.487
4.0532 | -33.9782 | -7.4975
4.0593 | -33.9786 | -7.5086
4.0654 | -33.9789 | -7.5191
4.0715 | -33.9792 | -7.5305
4.0776 | -33.9795 | -7.5412
4.0838 | -33.9799 | -7.5525
4.0899 | -33.9802 | -7.563
4.096 | -33.9805 | -7.5738
4.1021 | -33.9807 | -7.5844
4.1082 | -33.981 | -7.5951
4.1143 | -33.9813 | -7.6055
4.1204 | -33.9816 | -7.6161
4.1265 | -33.9819 | -7.6265
4.1326 | -33.9822 | -7.637
4.1387 | -33.9826 | -7.6476
4.1448 | -33.9829 | -7.6584
4.1509 | -33.9831 | -7.6693
4.157 | -33.9834 | -7.6805
4.1631 | -33.9837 | -7.6915
4.1692 | -33.984 | -7.7021
4.1753 | -33.9843 | -7.7128
4.1814 | -33.9845 | -7.7232
4.1875 | -33.9848 | -7.7342
4.1936 | -33.9851 | -7.7451
4.1997 | -33.9854 | -7.7555
4.2058 | -33.9857 | -7.7664
4.2119 | -33.9859 | -7.777
4.218 | -33.9862 | -7.7875
4.2241 | -33.9864 | -7.798
4.2303 | -33.9867 | -7.8085
4.2364 | -33.987 | -7.8192
4.2425 | -33.9872 | -7.83
4.2486 | -33.9874 | -7.8408
4.2547 | -33.9876 | -7.8512
4.2608 | -33.9878 | -7.862
4.2669 | -33.9881 | -7.8724
4.273 | -33.9883 | -7.8832
4.2791 | -33.9886 | -7.8939
4.2852 | -33.9888 | -7.9047
4.2913 | -33.989 | -7.9151
4.2974 | -33.9893 | -7.926
4.3035 | -33.9896 | -7.937
4.3096 | -33.9898 | -7.9478
4.3157 | -33.99 | -7.9583
4.3218 | -33.9903 | -7.9688
4.3279 | -33.9905 | -7.9795
4.334 | -33.9907 | -7.9901
4.3401 | -33.991 | -8.0003
4.3462 | -33.9913 | -8.011
4.3523 | -33.9915 | -8.0212
4.3584 | -33.9917 | -8.0314
4.3645 | -33.992 | -8.0418
4.3707 | -33.9922 | -8.0521
4.3768 | -33.9924 | -8.0627
4.3829 | -33.9927 | -8.0736
4.389 | -33.9929 | -8.0842
4.3951 | -33.9931 | -8.0945
4.4012 | -33.9933 | -8.1054
4.4073 | -33.9935 | -8.1158
4.4134 | -33.9937 | -8.1262
4.4195 | -33.994 | -8.1367
4.4256 | -33.9943 | -8.147
4.4317 | -33.9946 | -8.1575
4.4378 | -33.9948 | -8.1677
4.4439 | -33.995 | -8.1779
4.45 | -33.9951 | -8.1885
4.4561 | -33.9953 | -8.1989
4.4622 | -33.9955 | -8.2095
4.4683 | -33.9958 | -8.2199
4.4744 | -33.996 | -8.2307
4.4805 | -33.9962 | -8.2408
4.4866 | -33.9964 | -8.2512
4.4927 | -33.9967 | -8.2619
4.4988 | -33.9969 | -8.2723
4.5049 | -33.9971 | -8.2824
4.511 | -33.9974 | -8.2926
4.5172 | -33.9976 | -8.303
4.5233 | -33.9978 | -8.3139
4.5294 | -33.9981 | -8.3242
4.5355 | -33.9983 | -8.3346
4.5416 | -33.9985 | -8.345
4.5477 | -33.9987 | -8.3554
4.5538 | -33.9989 | -8.3654
4.5599 | -33.9992 | -8.3758
4.566 | -33.9994 | -8.3862
4.5721 | -33.9996 | -8.3967
4.5782 | -33.9999 | -8.4073
4.5843 | -34.0001 | -8.4178
4.5904 | -34.0004 | -8.4283
4.5965 | -34.0006 | -8.4386
4.6026 | -34.0008 | -8.4495
4.6087 | -34.001 | -8.4598
4.6148 | -34.0012 | -8.4705
4.6209 | -34.0015 | -8.4808
4.627 | -34.0018 | -8.4915
4.6331 | -34.0019 | -8.5018
4.6392 | -34.0021 | -8.5123
4.6453 | -34.0024 | -8.5227
4.6514 | -34.0026 | -8.5332
4.6576 | -34.0028 | -8.544
4.6637 | -34.003 | -8.5541
4.6698 | -34.0032 | -8.5643
4.6759 | -34.0035 | -8.5751
4.682 | -34.0037 | -8.5855
4.6881 | -34.0038 | -8.5959
4.6942 | -34.0041 | -8.6062
4.7003 | -34.0043 | -8.6167
4.7064 | -34.0045 | -8.6269
4.7125 | -34.0048 | -8.6376
4.7186 | -34.005 | -8.6478
4.7247 | -34.0052 | -8.6583
4.7308 | -34.0055 | -8.6686
4.7369 | -34.0057 | -8.6792
4.743 | -34.0059 | -8.69
4.7491 | -34.0062 | -8.7007
4.7552 | -34.0064 | -8.7112
4.7613 | -34.0065 | -8.7218
4.7674 | -34.0068 | -8.7321
4.7735 | -34.007 | -8.743
4.7796 | -34.0073 | -8.7536
4.7857 | -34.0075 | -8.7642
4.7918 | -34.0077 | -8.7745
4.7979 | -34.008 | -8.7848
4.8041 | -34.0082 | -8.7953
4.8102 | -34.0085 | -8.8054
4.8163 | -34.0087 | -8.8158
4.8224 | -34.0089 | -8.8263
4.8285 | -34.0092 | -8.8364
4.8346 | -34.0094 | -8.8467
4.8407 | -34.0096 | -8.8572
4.8468 | -34.0099 | -8.868
4.8529 | -34.0101 | -8.8787
4.859 | -34.0104 | -8.8895
4.8651 | -34.0106 | -8.8997
4.8712 | -34.0109 | -8.9099
4.8773 | -34.011 | -8.9204
4.8834 | -34.0113 | -8.9309
4.8895 | -34.0115 | -8.9412
4.8956 | -34.0117 | -8.9516
4.9017 | -34.0119 | -8.9619
4.9078 | -34.0122 | -8.9726
4.9139 | -34.0124 | -8.9827
4.92 | -34.0127 | -8.9928
4.9261 | -34.013 | -9.0029
4.9322 | -34.0132 | -9.0134
4.9383 | -34.0135 | -9.0238
4.9445 | -34.0138 | -9.0343
4.9506 | -34.014 | -9.0445
4.9567 | -34.0142 | -9.0547
4.9628 | -34.0145 | -9.0651
4.9689 | -34.0147 | -9.0759
4.975 | -34.0149 | -9.0866
4.9811 | -34.0152 | -9.0968
4.9872 | -34.0155 | -9.1071
4.9933 | -34.0158 | -9.1178
4.9994 | -34.0161 | -9.1286
5.0055 | -34.0164 | -9.1391
5.0116 | -34.0166 | -9.15
5.0177 | -34.017 | -9.1607
5.0238 | -34.0172 | -9.1715
5.0299 | -34.0176 | -9.1823
5.036 | -34.0178 | -9.1928
5.0421 | -34.0181 | -9.2032
5.0482 | -34.0183 | -9.2137
5.0543 | -34.0186 | -9.2239
5.0604 | -34.0189 | -9.2344
5.0665 | -34.0192 | -9.2446
5.0726 | -34.0195 | -9.2554
5.0787 | -34.0198 | -9.2661
5.0848 | -34.0201 | -9.2766
5.091 | -34.0203 | -9.2869
5.0971 | -34.0206 | -9.2971
5.1032 | -34.0209 | -9.3079
5.1093 | -34.0212 | -9.3184
5.1154 | -34.0214 | -9.3286
5.1215 | -34.0217 | -9.3389
5.1276 | -34.022 | -9.3491
5.1337 | -34.0223 | -9.3593
5.1398 | -34.0226 | -9.3695
5.1459 | -34.0229 | -9.3802
5.152 | -34.0232 | -9.3909
5.1581 | -34.0234 | -9.4016
5.1642 | -34.0238 | -9.4121
5.1703 | -34.024 | -9.4226
5.1764 | -34.0244 | -9.4329
5.1825 | -34.0246 | -9.4431
5.1886 | -34.025 | -9.4533
5.1947 | -34.0252 | -9.4638
5.2008 | -34.0255 | -9.4744
5.2069 | -34.0258 | -9.4851
5.213 | -34.0261 | -9.4953
5.2191 | -34.0265 | -9.5058
5.2252 | -34.0267 | -9.5164
5.2314 | -34.027 | -9.5271
5.2375 | -34.0273 | -9.5373
5.2436 | -34.0276 | -9.5479
5.2497 | -34.0279 | -9.5581
5.2558 | -34.0282 | -9.5689
5.2619 | -34.0285 | -9.5797
5.268 | -34.0289 | -9.5904
5.2741 | -34.0292 | -9.601
5.2802 | -34.0295 | -9.6119
5.2863 | -34.0298 | -9.6223
5.2924 | -34.0302 | -9.633
5.2985 | -34.0305 | -9.6436
5.3046 | -34.0308 | -9.6542
5.3107 | -34.0311 | -9.665
5.3168 | -34.0314 | -9.6756
5.3229 | -34.0317 | -9.6862
5.329 | -34.032 | -9.6968
5.3351 | -34.0324 | -9.707
5.3412 | -34.0326 | -9.7175
5.3473 | -34.033 | -9.7282
5.3534 | -34.0333 | -9.7388
5.3595 | -34.0336 | -9.7495
5.3656 | -34.034 | -9.76
5.3717 | -34.0343 | -9.7709
5.3779 | -34.0346 | -9.7818
5.384 | -34.0349 | -9.7924
5.3901 | -34.0353 | -9.8032
5.3962 | -34.0356 | -9.8136
5.4023 | -34.0359 | -9.8244
5.4084 | -34.0363 | -9.8349
5.4145 | -34.0366 | -9.8453
5.4206 | -34.0369 | -9.8559
5.4267 | -34.0372 | -9.8668
5.4328 | -34.0376 | -9.8775
5.4389 | -34.0379 | -9.8878
5.445 | -34.0382 | -9.8988
5.4511 | -34.0385 | -9.9092
5.4572 | -34.0389 | -9.9199
5.4633 | -34.0392 | -9.9305
5.4694 | -34.0395 | -9.9411
5.4755 | -34.0398 | -9.9515
5.4816 | -34.0402 | -9.9621
5.4877 | -34.0405 | -9.9727
5.4938 | -34.0408 | -9.983
5.4999 | -34.0412 | -9.9936
5.506 | -34.0416 | -10.0042
5.5121 | -34.0419 | -10.0147
5.5183 | -34.0422 | -10.0252
5.5244 | -34.0425 | -10.036
5.5305 | -34.0429 | -10.0465
5.5366 | -34.0432 | -10.0569
5.5427 | -34.0434 | -10.0677
5.5488 | -34.0438 | -10.0784
5.5549 | -34.0441 | -10.0891
5.561 | -34.0444 | -10.1
5.5671 | -34.0448 | -10.1104
5.5732 | -34.0451 | -10.1213
5.5793 | -34.0454 | -10.1323
5.5854 | -34.0458 | -10.1429
5.5915 | -34.0461 | -10.1539
5.5976 | -34.0465 | -10.1646
5.6037 | -34.0468 | -10.1753
5.6098 | -34.0471 | -10.1859
5.6159 | -34.0474 | -10.1963
5.622 | -34.0477 | -10.2071
5.6281 | -34.0481 | -10.2179
5.6342 | -34.0484 | -10.2289
5.6403 | -34.0488 | -10.2396
5.6464 | -34.0491 | -10.2504
5.6525 | -34.0494 | -10.2612
5.6586 | -34.0498 | -10.2715
5.6648 | -34.0501 | -10.2823
5.6709 | -34.0505 | -10.2931
5.677 | -34.0508 | -10.3038
5.6831 | -34.0512 | -10.3146
5.6892 | -34.0515 | -10.3254
5.6953 | -34.0518 | -10.3361
5.7014 | -34.0522 | -10.347
5.7075 | -34.0525 | -10.3576
5.7136 | -34.0528 | -10.3683
5.7197 | -34.0531 | -10.379
5.7258 | -34.0534 | -10.3895
5.7319 | -34.0538 | -10.4003
5.738 | -34.0541 | -10.411
5.7441 | -34.0545 | -10.4216
5.7502 | -34.0548 | -10.432
5.7563 | -34.0552 | -10.4427
5.7624 | -34.0555 | -10.4536
5.7685 | -34.0559 | -10.4644
5.7746 | -34.0562 | -10.4753
5.7807 | -34.0564 | -10.4861
5.7868 | -34.0568 | -10.497
5.7929 | -34.0571 | -10.5076
5.799 | -34.0574 | -10.5182
5.8052 | -34.0577 | -10.5287
5.8113 | -34.0581 | -10.5393
5.8174 | -34.0584 | -10.5498
5.8235 | -34.0587 | -10.5605
5.8296 | -34.059 | -10.5714
5.8357 | -34.0593 | -10.582
5.8418 | -34.0597 | -10.5928
5.8479 | -34.0599 | -10.6036
5.854 | -34.0603 | -10.6145
5.8601 | -34.0606 | -10.625
5.8662 | -34.0609 | -10.6358
5.8723 | -34.0612 | -10.6463
5.8784 | -34.0615 | -10.657
5.8845 | -34.0619 | -10.6676
5.8906 | -34.0622 | -10.6783
5.8967 | -34.0625 | -10.6888
5.9028 | -34.0628 | -10.6993
5.9089 | -34.0632 | -10.7098
5.915 | -34.0635 | -10.7203
5.9211 | -34.0638 | -10.7309
5.9272 | -34.0641 | -10.7413
5.9333 | -34.0645 | -10.7519
5.9394 | -34.0648 | -10.7626
5.9455 | -34.0651 | -10.773
5.9517 | -34.0655 | -10.7835
5.9578 | -34.0659 | -10.7943
5.9639 | -34.0662 | -10.8046
5.97 | -34.0666 | -10.8153
5.9761 | -34.0669 | -10.8259
5.9822 | -34.0672 | -10.8366
5.9883 | -34.0675 | -10.8472
5.9944 | -34.0679 | -10.8579
6.0005 | -34.0682 | -10.8687
6.0066 | -34.0685 | -10.8791
6.0127 | -34.0689 | -10.8895
6.0188 | -34.0692 | -10.9004
6.0249 | -34.0695 | -10.9111
6.031 | -34.0698 | -10.9216
6.0371 | -34.0702 | -10.9321
6.0432 | -34.0705 | -10.9426
6.0493 | -34.0708 | -10.9531
6.0554 | -34.0712 | -10.9635
6.0615 | -34.0715 | -10.9741
6.0676 | -34.0718 | -10.985
6.0737 | -34.0722 | -10.9957
6.0798 | -34.0725 | -11.0062
6.0859 | -34.0729 | -11.0167
6.0921 | -34.0732 | -11.0271
6.0982 | -34.0735 | -11.0376
6.1043 | -34.0739 | -11.0481
6.1104 | -34.0742 | -11.0589
6.1165 | -34.0745 | -11.069
6.1226 | -34.0749 | -11.0795
6.1287 | -34.0752 | -11.09
6.1348 | -34.0756 | -11.1006
6.1409 | -34.0759 | -11.1111
6.147 | -34.0762 | -11.1216
6.1531 | -34.0765 | -11.1321
6.1592 | -34.0768 | -11.1425
6.1653 | -34.0771 | -11.1529
6.1714 | -34.0774 | -11.1634
6.1775 | -34.0778 | -11.1741
6.1836 | -34.0781 | -11.1846
6.1897 | -34.0785 | -11.1951
6.1958 | -34.0788 | -11.2059
6.2019 | -34.0791 | -11.2165
6.208 | -34.0795 | -11.2271
6.2141 | -34.0798 | -11.2375
6.2202 | -34.0801 | -11.2481
6.2263 | -34.0805 | -11.2585
6.2325 | -34.0808 | -11.2687
6.2386 | -34.0811 | -11.2789
6.2447 | -34.0815 | -11.2897
6.2508 | -34.0818 | -11.3001
6.2569 | -34.0822 | -11.3105
6.263 | -34.0825 | -11.321
6.2691 | -34.0828 | -11.3314
6.2752 | -34.0831 | -11.342
6.2813 | -34.0834 | -11.3524
6.2874 | -34.0837 | -11.3627
6.2935 | -34.0841 | -11.373
6.2996 | -34.0844 | -11.3835
6.3057 | -34.0847 | -11.394
6.3118 | -34.085 | -11.4046
6.3179 | -34.0853 | -11.415
6.324 | -34.0856 | -11.4254
6.3301 | -34.086 | -11.4357
6.3362 | -34.0863 | -11.4461
6.3423 | -34.0866 | -11.4567
6.3484 | -34.087 | -11.4672
6.3545 | -34.0873 | -11.4773
6.3606 | -34.0876 | -11.4875
6.3667 | -34.0879 | -11.4977
6.3728 | -34.0882 | -11.5082
6.379 | -34.0886 | -11.5187
6.3851 | -34.0889 | -11.5288
6.3912 | -34.0892 | -11.5392
6.3973 | -34.0895 | -11.5496
6.4034 | -34.0898 | -11.5598
6.4095 | -34.0901 | -11.5701
6.4156 | -34.0904 | -11.5806
6.4217 | -34.0907 | -11.5909
6.4278 | -34.0911 | -11.6014
6.4339 | -34.0914 | -11.6117
6.44 | -34.0918 | -11.6219
6.4461 | -34.0921 | -11.6325
6.4522 | -34.0924 | -11.6426
6.4583 | -34.0927 | -11.6532
6.4644 | -34.093 | -11.6636
6.4705 | -34.0933 | -11.6736
6.4766 | -34.0936 | -11.684
6.4827 | -34.0939 | -11.6944
6.4888 | -34.0942 | -11.7046
6.4949 | -34.0946 | -11.7146
6.501 | -34.0949 | -11.7251
6.5071 | -34.0952 | -11.7351
6.5132 | -34.0955 | -11.7455
6.5194 | -34.0958 | -11.7559
6.5255 | -34.0962 | -11.766
6.5316 | -34.0965 | -11.7765
6.5377 | -34.0968 | -11.7868
6.5438 | -34.0971 | -11.7968
6.5499 | -34.0974 | -11.807
6.556 | -34.0977 | -11.8173
6.5621 | -34.098 | -11.8279
6.5682 | -34.0983 | -11.8383
6.5743 | -34.0987 | -11.8484
6.5804 | -34.099 | -11.8588
6.5865 | -34.0994 | -11.8691
6.5926 | -34.0997 | -11.8793
6.5987 | -34.1 | -11.8893
6.6048 | -34.1003 | -11.8996
6.6109 | -34.1007 | -11.9099
6.617 | -34.1009 | -11.9203
6.6231 | -34.1012 | -11.9305
6.6292 | -34.1016 | -11.9407
6.6353 | -34.1019 | -11.9513
6.6414 | -34.1022 | -11.9615
6.6475 | -34.1025 | -11.9716
6.6536 | -34.1028 | -11.9817
6.6597 | -34.1032 | -11.9917
6.6659 | -34.1035 | -12.002
6.672 | -34.1038 | -12.0121
6.6781 | -34.1041 | -12.0224
6.6842 | -34.1045 | -12.0324
6.6903 | -34.1048 | -12.0427
6.6964 | -34.1052 | -12.0529
6.7025 | -34.1055 | -12.063
6.7086 | -34.1058 | -12.0731
6.7147 | -34.1061 | -12.0835
6.7208 | -34.1064 | -12.0938
6.7269 | -34.1067 | -12.104
6.733 | -34.107 | -12.114
6.7391 | -34.1074 | -12.1242
6.7452 | -34.1077 | -12.1342
6.7513 | -34.108 | -12.1441
6.7574 | -34.1083 | -12.1544
6.7635 | -34.1087 | -12.1644
6.7696 | -34.109 | -12.1744
6.7757 | -34.1094 | -12.1847
6.7818 | -34.1097 | -12.1948
6.7879 | -34.11 | -12.2048
6.794 | -34.1104 | -12.2148
6.8001 | -34.1106 | -12.2249
6.8063 | -34.111 | -12.2351
6.8124 | -34.1113 | -12.2453
6.8185 | -34.1116 | -12.2556
6.8246 | -34.1119 | -12.2656
6.8307 | -34.1122 | -12.2756
6.8368 | -34.1125 | -12.286
6.8429 | -34.1129 | -12.2962
6.849 | -34.1132 | -12.3064
6.8551 | -34.1136 | -12.3164
6.8612 | -34.1139 | -12.3265
6.8673 | -34.1142 | -12.3365
6.8734 | -34.1145 | -12.3467
6.8795 | -34.1148 | -12.3569
6.8856 | -34.1152 | -12.367
6.8917 | -34.1155 | -12.3771
6.8978 | -34.1158 | -12.387
6.9039 | -34.1162 | -12.3969
6.91 | -34.1165 | -12.407
6.9161 | -34.1169 | -12.4169
6.9222 | -34.1172 | -12.427
6.9283 | -34.1176 | -12.4368
6.9344 | -34.1179 | -12.447
6.9405 | -34.1183 | -12.4571
6.9466 | -34.1186 | -12.4674
6.9528 | -34.119 | -12.4777
6.9589 | -34.1193 | -12.4879
6.965 | -34.1197 | -12.498
6.9711 | -34.12 | -12.5082
6.9772 | -34.1204 | -12.5183
6.9833 | -34.1207 | -12.5283
6.9894 | -34.1211 | -12.5387
6.9955 | -34.1214 | -12.5485
7.0016 | -34.1218 | -12.5586
7.0077 | -34.1221 | -12.5687
7.0138 | -34.1225 | -12.5787
7.0199 | -34.1228 | -12.5887
7.026 | -34.1231 | -12.5987
7.0321 | -34.1235 | -12.6087
7.0382 | -34.1238 | -12.6184
7.0443 | -34.1242 | -12.6287
7.0504 | -34.1245 | -12.6389
7.0565 | -34.1248 | -12.649
7.0626 | -34.1252 | -12.6587
7.0687 | -34.1256 | -12.6684
7.0748 | -34.1259 | -12.6782
7.0809 | -34.1263 | -12.6882
7.087 | -34.1266 | -12.6984
7.0932 | -34.1269 | -12.7084
7.0993 | -34.1273 | -12.7184
7.1054 | -34.1276 | -12.7283
7.1115 | -34.128 | -12.7382
7.1176 | -34.1283 | -12.7481
7.1237 | -34.1287 | -12.7581
7.1298 | -34.129 | -12.768
7.1359 | -34.1293 | -12.778
7.142 | -34.1297 | -12.7878
7.1481 | -34.13 | -12.7977
7.1542 | -34.1304 | -12.8078
7.1603 | -34.1307 | -12.8177
7.1664 | -34.1311 | -12.8277
7.1725 | -34.1315 | -12.8378
7.1786 | -34.1319 | -12.8478
7.1847 | -34.1322 | -12.8578
7.1908 | -34.1326 | -12.868
7.1969 | -34.1329 | -12.8782
7.203 | -34.1332 | -12.8881
7.2091 | -34.1336 | -12.898
7.2152 | -34.134 | -12.908
7.2213 | -34.1343 | -12.9178
7.2274 | -34.1347 | -12.9277
7.2335 | -34.135 | -12.9377
7.2397 | -34.1354 | -12.9476
7.2458 | -34.1358 | -12.9574
7.2519 | -34.1362 | -12.9674
7.258 | -34.1365 | -12.9775
7.2641 | -34.1369 | -12.9875
7.2702 | -34.1372 | -12.9977
7.2763 | -34.1376 | -13.0077
7.2824 | -34.138 | -13.0178
7.2885 | -34.1383 | -13.0278
7.2946 | -34.1387 | -13.0379
7.3007 | -34.139 | -13.048
7.3068 | -34.1393 | -13.0583
7.3129 | -34.1397 | -13.0685
7.319 | -34.1401 | -13.0784
7.3251 | -34.1404 | -13.0882
7.3312 | -34.1408 | -13.0983
7.3373 | -34.1411 | -13.1081
7.3434 | -34.1415 | -13.1179
7.3495 | -34.1419 | -13.1278
7.3556 | -34.1423 | -13.1378
7.3617 | -34.1426 | -13.148
7.3678 | -34.143 | -13.1579
7.3739 | -34.1433 | -13.1681
7.3801 | -34.1437 | -13.1783
7.3862 | -34.1441 | -13.1883
7.3923 | -34.1445 | -13.1985
7.3984 | -34.1448 | -13.2084
7.4045 | -34.1452 | -13.2184
7.4106 | -34.1456 | -13.2283
7.4167 | -34.1459 | -13.2385
7.4228 | -34.1463 | -13.2486
7.4289 | -34.1467 | -13.2587
7.435 | -34.147 | -13.2685
7.4411 | -34.1474 | -13.2786
7.4472 | -34.1477 | -13.2887
7.4533 | -34.1481 | -13.2985
7.4594 | -34.1485 | -13.3088
7.4655 | -34.1489 | -13.3185
7.4716 | -34.1493 | -13.3287
7.4777 | -34.1496 | -13.3387
7.4838 | -34.15 | -13.3486
7.4899 | -34.1503 | -13.3586
7.496 | -34.1507 | -13.3685
7.5021 | -34.1511 | -13.3784
7.5082 | -34.1515 | -13.3884
7.5143 | -34.1518 | -13.3984
7.5204 | -34.1522 | -13.4085
7.5266 | -34.1526 | -13.4184
7.5327 | -34.1529 | -13.4283
7.5388 | -34.1533 | -13.4387
7.5449 | -34.1537 | -13.4489
7.551 | -34.1541 | -13.4589
7.5571 | -34.1545 | -13.469
7.5632 | -34.1549 | -13.479
7.5693 | -34.1553 | -13.489
7.5754 | -34.1556 | -13.499
7.5815 | -34.156 | -13.5091
7.5876 | -34.1563 | -13.5193
7.5937 | -34.1568 | -13.5296
7.5998 | -34.1572 | -13.5396
7.6059 | -34.1576 | -13.5498
7.612 | -34.1579 | -13.5599
7.6181 | -34.1583 | -13.5699
7.6242 | -34.1587 | -13.5802
7.6303 | -34.1591 | -13.5907
7.6364 | -34.1594 | -13.6008
7.6425 | -34.1598 | -13.6108
7.6486 | -34.1601 | -13.6209
7.6547 | -34.1605 | -13.6308
7.6608 | -34.1609 | -13.6409
7.667 | -34.1613 | -13.6511
7.6731 | -34.1617 | -13.6613
7.6792 | -34.1621 | -13.6716
7.6853 | -34.1624 | -13.6818
7.6914 | -34.1628 | -13.6919
7.6975 | -34.1632 | -13.7021
7.7036 | -34.1636 | -13.7123
7.7097 | -34.164 | -13.7225
7.7158 | -34.1644 | -13.7327
7.7219 | -34.1647 | -13.7426
7.728 | -34.1651 | -13.7528
7.7341 | -34.1655 | -13.7631
7.7402 | -34.1659 | -13.7732
7.7463 | -34.1663 | -13.7834
7.7524 | -34.1666 | -13.7937
7.7585 | -34.167 | -13.8038
7.7646 | -34.1674 | -13.814
7.7707 | -34.1678 | -13.8242
7.7768 | -34.1681 | -13.8345
7.7829 | -34.1685 | -13.845
7.789 | -34.1689 | -13.8553
7.7951 | -34.1693 | -13.8654
7.8012 | -34.1696 | -13.8756
7.8073 | -34.1701 | -13.886
7.8135 | -34.1704 | -13.8961
7.8196 | -34.1709 | -13.9065
7.8257 | -34.1712 | -13.9169
7.8318 | -34.1716 | -13.9271
7.8379 | -34.172 | -13.9372
7.844 | -34.1724 | -13.9473
7.8501 | -34.1727 | -13.9576
7.8562 | -34.1731 | -13.968
7.8623 | -34.1735 | -13.9781
7.8684 | -34.1739 | -13.9884
7.8745 | -34.1743 | -13.9987
7.8806 | -34.1747 | -14.0091
7.8867 | -34.1751 | -14.0191
7.8928 | -34.1754 | -14.029
7.8989 | -34.1758 | -14.0393
7.905 | -34.1762 | -14.0495
7.9111 | -34.1766 | -14.0599
7.9172 | -34.177 | -14.0701
7.9233 | -34.1773 | -14.0803
7.9294 | -34.1777 | -14.0906
7.9355 | -34.1781 | -14.1006
7.9416 | -34.1784 | -14.1107
7.9477 | -34.1788 | -14.1208
7.9539 | -34.1792 | -14.1309
7.96 | -34.1796 | -14.1412
7.9661 | -34.18 | -14.1516
7.9722 | -34.1804 | -14.1616
7.9783 | -34.1808 | -14.1721
7.9844 | -34.1812 | -14.1823
7.9905 | -34.1816 | -14.1927
7.9966 | -34.1819 | -14.203
8.0027 | -34.1823 | -14.2132
8.0088 | -34.1827 | -14.2234
8.0149 | -34.183 | -14.2336
8.021 | -34.1834 | -14.2439
8.0271 | -34.1838 | -14.2544
8.0332 | -34.1842 | -14.2646
8.0393 | -34.1846 | -14.2749
8.0454 | -34.185 | -14.2855
8.0515 | -34.1854 | -14.2958
8.0576 | -34.1858 | -14.3062
8.0637 | -34.1862 | -14.3164
8.0698 | -34.1865 | -14.3265
8.0759 | -34.1869 | -14.3366
8.082 | -34.1873 | -14.3468
8.0881 | -34.1877 | -14.3571
8.0942 | -34.1881 | -14.3674
8.1004 | -34.1885 | -14.3777
8.1065 | -34.1889 | -14.3877
8.1126 | -34.1892 | -14.398
8.1187 | -34.1896 | -14.4083
8.1248 | -34.19 | -14.4184
8.1309 | -34.1904 | -14.4284
8.137 | -34.1908 | -14.4384
8.1431 | -34.1911 | -14.4488
8.1492 | -34.1915 | -14.459
8.1553 | -34.1919 | -14.4693
8.1614 | -34.1923 | -14.4796
8.1675 | -34.1926 | -14.4898
8.1736 | -34.193 | -14.5003
8.1797 | -34.1934 | -14.5107
8.1858 | -34.1938 | -14.521
8.1919 | -34.1942 | -14.531
8.198 | -34.1945 | -14.5411
8.2041 | -34.195 | -14.5514
8.2102 | -34.1953 | -14.5615
8.2163 | -34.1957 | -14.5717
8.2224 | -34.1961 | -14.5819
8.2285 | -34.1965 | -14.5918
8.2346 | -34.1969 | -14.602
8.2408 | -34.1973 | -14.612
8.2469 | -34.1977 | -14.6224
8.253 | -34.1981 | -14.6325
8.2591 | -34.1985 | -14.6428
8.2652 | -34.1989 | -14.6531
8.2713 | -34.1993 | -14.6633
8.2774 | -34.1997 | -14.6737
8.2835 | -34.2001 | -14.6842
8.2896 | -34.2005 | -14.6942
8.2957 | -34.2009 | -14.7045
8.3018 | -34.2012 | -14.7147
8.3079 | -34.2016 | -14.7247
8.314 | -34.202 | -14.7349
8.3201 | -34.2024 | -14.7451
8.3262 | -34.2028 | -14.7553
8.3323 | -34.2032 | -14.7653
8.3384 | -34.2036 | -14.7757
8.3445 | -34.204 | -14.7859
8.3506 | -34.2044 | -14.7962
8.3567 | -34.2048 | -14.8065
8.3628 | -34.2052 | -14.8167
8.3689 | -34.2056 | -14.8268
8.375 | -34.206 | -14.8371
8.3812 | -34.2064 | -14.8473
8.3873 | -34.2068 | -14.8574
8.3934 | -34.2072 | -14.8678
8.3995 | -34.2076 | -14.8778
8.4056 | -34.208 | -14.8881
8.4117 | -34.2084 | -14.8984
8.4178 | -34.2087 | -14.9086
8.4239 | -34.2091 | -14.9186
8.43 | -34.2095 | -14.9288
8.4361 | -34.2099 | -14.9391
8.4422 | -34.2103 | -14.9491
8.4483 | -34.2107 | -14.9592
8.4544 | -34.2111 | -14.9694
8.4605 | -34.2115 | -14.9794
8.4666 | -34.2119 | -14.9898
8.4727 | -34.2123 | -14.9999
8.4788 | -34.2127 | -15.0101
8.4849 | -34.2131 | -15.0203
8.491 | -34.2134 | -15.0305
8.4971 | -34.2138 | -15.0404
8.5032 | -34.2142 | -15.0504
8.5093 | -34.2146 | -15.0603
8.5154 | -34.215 | -15.0706
8.5215 | -34.2154 | -15.0806
8.5277 | -34.2158 | -15.0907
8.5338 | -34.2162 | -15.1009
8.5399 | -34.2166 | -15.111
8.546 | -34.217 | -15.1209
8.5521 | -34.2174 | -15.1311
8.5582 | -34.2178 | -15.1411
8.5643 | -34.2182 | -15.1509
8.5704 | -34.2186 | -15.1607
8.5765 | -34.219 | -15.1707
8.5826 | -34.2194 | -15.1804
8.5887 | -34.2198 | -15.1905
8.5948 | -34.2202 | -15.2006
8.6009 | -34.2206 | -15.2105
8.607 | -34.221 | -15.2208
8.6131 | -34.2214 | -15.2309
8.6192 | -34.2218 | -15.2408
8.6253 | -34.2222 | -15.2508
8.6314 | -34.2226 | -15.2608
8.6375 | -34.2229 | -15.271
8.6436 | -34.2234 | -15.2811
8.6497 | -34.2237 | -15.291
8.6558 | -34.2241 | -15.3011
8.6619 | -34.2244 | -15.3113
8.6681 | -34.2249 | -15.3214
8.6742 | -34.2253 | -15.3313
8.6803 | -34.2257 | -15.3413
8.6864 | -34.2261 | -15.351
8.6925 | -34.2264 | -15.3608
8.6986 | -34.2268 | -15.3709
8.7047 | -34.2272 | -15.3812
8.7108 | -34.2276 | -15.3912
8.7169 | -34.228 | -15.4012
8.723 | -34.2284 | -15.4111
8.7291 | -34.2288 | -15.4209
8.7352 | -34.2292 | -15.4311
8.7413 | -34.2296 | -15.4412
8.7474 | -34.23 | -15.451
8.7535 | -34.2304 | -15.4611
8.7596 | -34.2308 | -15.4709
8.7657 | -34.2312 | -15.481
8.7718 | -34.2316 | -15.4909
8.7779 | -34.232 | -15.501
8.784 | -34.2324 | -15.511
8.7901 | -34.2328 | -15.5211
8.7962 | -34.2332 | -15.5309
8.8023 | -34.2336 | -15.5409
8.8084 | -34.234 | -15.5509
8.8146 | -34.2344 | -15.5608
8.8207 | -34.2347 | -15.5708
8.8268 | -34.2352 | -15.5804
8.8329 | -34.2355 | -15.5903
8.839 | -34.236 | -15.6004
8.8451 | -34.2364 | -15.6103
8.8512 | -34.2368 | -15.6201
8.8573 | -34.2371 | -15.6301
8.8634 | -34.2375 | -15.6397
8.8695 | -34.238 | -15.6499
8.8756 | -34.2383 | -15.6598
8.8817 | -34.2387 | -15.6696
8.8878 | -34.2391 | -15.6797
8.8939 | -34.2395 | -15.6896
8.9 | -34.2399 | -15.6995
8.9061 | -34.2403 | -15.7094
8.9122 | -34.2407 | -15.7192
8.9183 | -34.2411 | -15.7292
8.9244 | -34.2415 | -15.7393
8.9305 | -34.242 | -15.7491
8.9366 | -34.2424 | -15.759
8.9427 | -34.2428 | -15.7691
8.9488 | -34.2432 | -15.779
8.955 | -34.2436 | -15.7888
8.9611 | -34.244 | -15.7988
8.9672 | -34.2445 | -15.8085
8.9733 | -34.2449 | -15.8183
8.9794 | -34.2453 | -15.828
8.9855 | -34.2457 | -15.838
8.9916 | -34.2461 | -15.848
8.9977 | -34.2465 | -15.858
9.0038 | -34.247 | -15.8677
9.0099 | -34.2474 | -15.8776
9.016 | -34.2478 | -15.8876
9.0221 | -34.2482 | -15.8972
9.0282 | -34.2486 | -15.907
9.0343 | -34.249 | -15.9168
9.0404 | -34.2494 | -15.9267
9.0465 | -34.2499 | -15.9368
9.0526 | -34.2503 | -15.9465
9.0587 | -34.2507 | -15.9565
9.0648 | -34.2511 | -15.9662
9.0709 | -34.2515 | -15.9761
9.077 | -34.2519 | -15.986
9.0831 | -34.2524 | -15.996
9.0892 | -34.2528 | -16.006
9.0953 | -34.2532 | -16.0159
9.1015 | -34.2536 | -16.0257
9.1076 | -34.2541 | -16.0354
9.1137 | -34.2545 | -16.0453
9.1198 | -34.2549 | -16.0552
9.1259 | -34.2553 | -16.065
9.132 | -34.2557 | -16.075
9.1381 | -34.2562 | -16.0846
9.1442 | -34.2566 | -16.094
9.1503 | -34.257 | -16.1041
9.1564 | -34.2575 | -16.1138
9.1625 | -34.2579 | -16.1236
9.1686 | -34.2583 | -16.1335
9.1747 | -34.2587 | -16.1434
9.1808 | -34.2591 | -16.1531
9.1869 | -34.2596 | -16.1628
9.193 | -34.26 | -16.1727
9.1991 | -34.2604 | -16.1826
9.2052 | -34.2608 | -16.1925
9.2113 | -34.2613 | -16.2023
9.2174 | -34.2617 | -16.212
9.2235 | -34.2622 | -16.2218
9.2296 | -34.2626 | -16.2317
9.2357 | -34.263 | -16.2417
9.2419 | -34.2635 | -16.2513
9.248 | -34.264 | -16.2611
9.2541 | -34.2644 | -16.271
9.2602 | -34.2648 | -16.2809
9.2663 | -34.2652 | -16.2904
9.2724 | -34.2657 | -16.3002
9.2785 | -34.2661 | -16.3099
9.2846 | -34.2665 | -16.3197
9.2907 | -34.2669 | -16.3295
9.2968 | -34.2674 | -16.3394
9.3029 | -34.2678 | -16.3494
9.309 | -34.2682 | -16.3591
9.3151 | -34.2686 | -16.3688
9.3212 | -34.2691 | -16.3783
9.3273 | -34.2695 | -16.3879
9.3334 | -34.27 | -16.3976
9.3395 | -34.2704 | -16.4073
9.3456 | -34.2708 | -16.4169
9.3517 | -34.2712 | -16.4266
9.3578 | -34.2716 | -16.4363
9.3639 | -34.2721 | -16.4461
9.37 | -34.2725 | -16.4557
9.3761 | -34.273 | -16.4658
9.3822 | -34.2734 | -16.4756
9.3884 | -34.2739 | -16.4853
9.3945 | -34.2743 | -16.4953
9.4006 | -34.2748 | -16.5051
9.4067 | -34.2752 | -16.5148
9.4128 | -34.2757 | -16.5248
9.4189 | -34.2761 | -16.5348
9.425 | -34.2765 | -16.5445
9.4311 | -34.277 | -16.5543
9.4372 | -34.2774 | -16.5641
9.4433 | -34.2779 | -16.5738
9.4494 | -34.2783 | -16.5834
9.4555 | -34.2787 | -16.5933
9.4616 | -34.2792 | -16.6031
9.4677 | -34.2796 | -16.6128
9.4738 | -34.28 | -16.6225
9.4799 | -34.2805 | -16.6322
9.486 | -34.2809 | -16.6419
9.4921 | -34.2813 | -16.6516
9.4982 | -34.2818 | -16.6611
9.5043 | -34.2822 | -16.6708
9.5104 | -34.2826 | -16.6804
9.5165 | -34.2831 | -16.6902
9.5226 | -34.2835 | -16.7001
9.5288 | -34.2839 | -16.7096
9.5349 | -34.2843 | -16.7195
9.541 | -34.2847 | -16.7291
9.5471 | -34.2852 | -16.7387
9.5532 | -34.2856 | -16.7485
9.5593 | -34.2861 | -16.7583
9.5654 | -34.2865 | -16.7679
9.5715 | -34.287 | -16.7776
9.5776 | -34.2874 | -16.7874
9.5837 | -34.2878 | -16.7971
9.5898 | -34.2882 | -16.8069
9.5959 | -34.2887 | -16.8166
9.602 | -34.2891 | -16.8261
9.6081 | -34.2896 | -16.8358
9.6142 | -34.29 | -16.8454
9.6203 | -34.2905 | -16.8551
9.6264 | -34.2909 | -16.8647
9.6325 | -34.2913 | -16.8745
9.6386 | -34.2918 | -16.8841
9.6447 | -34.2922 | -16.8937
9.6508 | -34.2926 | -16.9034
9.6569 | -34.293 | -16.9132
9.663 | -34.2935 | -16.9229
9.6691 | -34.2939 | -16.9326
9.6753 | -34.2944 | -16.9423
9.6814 | -34.2948 | -16.9517
9.6875 | -34.2952 | -16.9614
9.6936 | -34.2956 | -16.9709
9.6997 | -34.2961 | -16.9808
9.7058 | -34.2966 | -16.9904
9.7119 | -34.297 | -17.0001
9.718 | -34.2975 | -17.01
9.7241 | -34.2979 | -17.0196
9.7302 | -34.2984 | -17.0293
9.7363 | -34.2988 | -17.0387
9.7424 | -34.2993 | -17.0484
9.7485 | -34.2997 | -17.0583
9.7546 | -34.3001 | -17.0678
9.7607 | -34.3005 | -17.0778
9.7668 | -34.301 | -17.0875
9.7729 | -34.3014 | -17.097
9.779 | -34.3019 | -17.1066
9.7851 | -34.3024 | -17.1164
9.7912 | -34.3029 | -17.1259
9.7973 | -34.3033 | -17.1359
9.8034 | -34.3038 | -17.1454
9.8095 | -34.3042 | -17.1554
9.8157 | -34.3047 | -17.1653
9.8218 | -34.3051 | -17.1751
9.8279 | -34.3055 | -17.1845
9.834 | -34.3059 | -17.1941
9.8401 | -34.3063 | -17.2037
9.8462 | -34.3068 | -17.213
9.8523 | -34.3072 | -17.2227
9.8584 | -34.3077 | -17.2323
9.8645 | -34.3081 | -17.2422
9.8706 | -34.3085 | -17.2519
9.8767 | -34.3089 | -17.2614
9.8828 | -34.3093 | -17.2711
9.8889 | -34.3098 | -17.2807
9.895 | -34.3102 | -17.2904
9.9011 | -34.3107 | -17.3001
9.9072 | -34.3111 | -17.3098
9.9133 | -34.3115 | -17.3195
9.9194 | -34.312 | -17.3289
9.9255 | -34.3124 | -17.3385
9.9316 | -34.3129 | -17.3482
9.9377 | -34.3133 | -17.3578
9.9438 | -34.3137 | -17.3675
9.9499 | -34.3142 | -17.3774
9.956 | -34.3146 | -17.3869
9.9622 | -34.3151 | -17.3966
9.9683 | -34.3155 | -17.4062
9.9744 | -34.316 | -17.4159
9.9805 | -34.3164 | -17.4255
9.9866 | -34.3168 | -17.4351
9.9927 | -34.3173 | -17.4445
9.9988 | -34.3177 | -17.4542
10.0049 | -34.3182 | -17.4638
10.011 | -34.3186 | -17.4736
10.0171 | -34.3191 | -17.4831
10.0232 | -34.3195 | -17.4925
10.0293 | -34.32 | -17.5022
10.0354 | -34.3205 | -17.5116
10.0415 | -34.3209 | -17.5215
10.0476 | -34.3213 | -17.5312
10.0537 | -34.3218 | -17.5411
10.0598 | -34.3222 | -17.5509
10.0659 | -34.3226 | -17.5604
10.072 | -34.323 | -17.5697
10.0781 | -34.3235 | -17.5793
10.0842 | -34.3239 | -17.5889
10.0903 | -34.3243 | -17.5985
10.0964 | -34.3248 | -17.6077
10.1026 | -34.3252 | -17.6174
10.1087 | -34.3256 | -17.627
10.1148 | -34.3261 | -17.6367
10.1209 | -34.3265 | -17.6462
10.127 | -34.3269 | -17.6556
10.1331 | -34.3274 | -17.6651
10.1392 | -34.3278 | -17.6745
10.1453 | -34.3283 | -17.6839
10.1514 | -34.3287 | -17.6936
10.1575 | -34.3291 | -17.7033
10.1636 | -34.3295 | -17.7126
10.1697 | -34.33 | -17.7224
10.1758 | -34.3304 | -17.7318
10.1819 | -34.3309 | -17.7414
10.188 | -34.3313 | -17.751
10.1941 | -34.3317 | -17.7604
10.2002 | -34.3321 | -17.7702
10.2063 | -34.3325 | -17.7799
10.2124 | -34.333 | -17.7895
10.2185 | -34.3334 | -17.7992
10.2246 | -34.3339 | -17.8086
10.2307 | -34.3343 | -17.8183
10.2368 | -34.3347 | -17.8277
10.2429 | -34.3351 | -17.8372
10.2491 | -34.3356 | -17.8467
10.2552 | -34.336 | -17.8565
10.2613 | -34.3364 | -17.8659
10.2674 | -34.3369 | -17.8753
10.2735 | -34.3373 | -17.8849
10.2796 | -34.3377 | -17.8946
10.2857 | -34.3382 | -17.9043
10.2918 | -34.3386 | -17.9139
10.2979 | -34.339 | -17.9235
10.304 | -34.3395 | -17.933
10.3101 | -34.3399 | -17.9426
10.3162 | -34.3403 | -17.9524
10.3223 | -34.3408 | -17.9619
10.3284 | -34.3412 | -17.9715
10.3345 | -34.3416 | -17.981
10.3406 | -34.3421 | -17.9905
10.3467 | -34.3425 | -18
10.3528 | -34.3429 | -18.0095
10.3589 | -34.3433 | -18.0189
10.365 | -34.3438 | -18.0287
10.3711 | -34.3442 | -18.0383
10.3772 | -34.3447 | -18.0477
10.3833 | -34.3451 | -18.0573
10.3895 | -34.3455 | -18.0668
10.3956 | -34.3459 | -18.0764
10.4017 | -34.3464 | -18.0858
10.4078 | -34.3468 | -18.0953
10.4139 | -34.3473 | -18.1052
10.42 | -34.3477 | -18.1149
10.4261 | -34.3481 | -18.1244
10.4322 | -34.3486 | -18.1338
10.4383 | -34.3489 | -18.1433
10.4444 | -34.3494 | -18.1526
10.4505 | -34.3498 | -18.1622
10.4566 | -34.3502 | -18.1717
10.4627 | -34.3507 | -18.1812
10.4688 | -34.3511 | -18.1906
10.4749 | -34.3516 | -18.2003
10.481 | -34.352 | -18.2098
10.4871 | -34.3525 | -18.2189
10.4932 | -34.3529 | -18.2283
10.4993 | -34.3533 | -18.2378
10.5054 | -34.3537 | -18.2473
10.5115 | -34.3541 | -18.257
10.5176 | -34.3546 | -18.2663
10.5237 | -34.355 | -18.2759
10.5298 | -34.3554 | -18.285
10.536 | -34.3558 | -18.2947
10.5421 | -34.3563 | -18.3041
10.5482 | -34.3567 | -18.3137
10.5543 | -34.3572 | -18.3232
10.5604 | -34.3576 | -18.3325
10.5665 | -34.358 | -18.3418
10.5726 | -34.3584 | -18.3514
10.5787 | -34.3589 | -18.3609
10.5848 | -34.3593 | -18.3701
10.5909 | -34.3597 | -18.3797
10.597 | -34.3602 | -18.3894
10.6031 | -34.3606 | -18.3991
10.6092 | -34.361 | -18.4087
10.6153 | -34.3615 | -18.4184
10.6214 | -34.3619 | -18.4278
10.6275 | -34.3624 | -18.4374
10.6336 | -34.3628 | -18.4471
10.6397 | -34.3632 | -18.4565
10.6458 | -34.3637 | -18.4661
10.6519 | -34.3641 | -18.4757
10.658 | -34.3645 | -18.4855
10.6641 | -34.3649 | -18.4951
10.6702 | -34.3654 | -18.5046
10.6764 | -34.3658 | -18.5144
10.6825 | -34.3663 | -18.5237
10.6886 | -34.3667 | -18.5332
10.6947 | -34.3671 | -18.5427
10.7008 | -34.3676 | -18.5521
10.7069 | -34.368 | -18.5616
10.713 | -34.3685 | -18.5711
10.7191 | -34.3689 | -18.5806
10.7252 | -34.3693 | -18.5899
10.7313 | -34.3697 | -18.5994
10.7374 | -34.3701 | -18.6089
10.7435 | -34.3706 | -18.6184
10.7496 | -34.371 | -18.6277
10.7557 | -34.3714 | -18.6374
10.7618 | -34.3718 | -18.6464
10.7679 | -34.3723 | -18.6559
10.774 | -34.3727 | -18.6653
10.7801 | -34.3732 | -18.675
10.7862 | -34.3736 | -18.6843
10.7923 | -34.374 | -18.694
10.7984 | -34.3744 | -18.7033
10.8045 | -34.3749 | -18.7127
10.8106 | -34.3753 | -18.7222
10.8168 | -34.3757 | -18.7315
10.8229 | -34.3762 | -18.7411
10.829 | -34.3766 | -18.7505
10.8351 | -34.3771 | -18.7597
10.8412 | -34.3775 | -18.7689
10.8473 | -34.378 | -18.7782
10.8534 | -34.3784 | -18.7875
10.8595 | -34.3788 | -18.7971
10.8656 | -34.3793 | -18.8064
10.8717 | -34.3797 | -18.8158
10.8778 | -34.3802 | -18.8253
10.8839 | -34.3806 | -18.8347
10.89 | -34.381 | -18.8441
10.8961 | -34.3815 | -18.8537
10.9022 | -34.3819 | -18.8632
10.9083 | -34.3823 | -18.8727
10.9144 | -34.3827 | -18.8822
10.9205 | -34.3832 | -18.8917
10.9266 | -34.3836 | -18.9012
10.9327 | -34.384 | -18.9106
10.9388 | -34.3845 | -18.9198
10.9449 | -34.3849 | -18.9292
10.951 | -34.3854 | -18.9386
10.9571 | -34.3858 | -18.9481
10.9633 | -34.3862 | -18.9577
10.9694 | -34.3867 | -18.9672
10.9755 | -34.3871 | -18.9768
10.9816 | -34.3875 | -18.9864
10.9877 | -34.388 | -18.9956
10.9938 | -34.3884 | -19.0051
10.9999 | -34.3888 | -19.0144
11.006 | -34.3893 | -19.0236
11.0121 | -34.3897 | -19.0331
11.0182 | -34.3901 | -19.0425
11.0243 | -34.3905 | -19.052
11.0304 | -34.391 | -19.0615
11.0365 | -34.3914 | -19.0707
11.0426 | -34.3919 | -19.08
11.0487 | -34.3923 | -19.0896
11.0548 | -34.3927 | -19.0989
11.0609 | -34.3932 | -19.1083
11.067 | -34.3936 | -19.1178
11.0731 | -34.3941 | -19.1272
11.0792 | -34.3946 | -19.1365
11.0853 | -34.395 | -19.146
11.0914 | -34.3955 | -19.1554
11.0975 | -34.3959 | -19.1648
11.1037 | -34.3963 | -19.1741
11.1098 | -34.3968 | -19.1835
11.1159 | -34.3972 | -19.1927
11.122 | -34.3976 | -19.2019
11.1281 | -34.3981 | -19.2113
11.1342 | -34.3985 | -19.2208
11.1403 | -34.399 | -19.2303
11.1464 | -34.3994 | -19.2396
11.1525 | -34.3998 | -19.2489
11.1586 | -34.4002 | -19.2583
11.1647 | -34.4007 | -19.2677
11.1708 | -34.4011 | -19.2771
11.1769 | -34.4015 | -19.2865
11.183 | -34.402 | -19.2957
11.1891 | -34.4025 | -19.3052
11.1952 | -34.4029 | -19.3145
11.2013 | -34.4034 | -19.324
11.2074 | -34.4038 | -19.3333
11.2135 | -34.4043 | -19.3429
11.2196 | -34.4047 | -19.3524
11.2257 | -34.4051 | -19.3619
11.2318 | -34.4056 | -19.3711
11.2379 | -34.406 | -19.3808
11.244 | -34.4065 | -19.3902
11.2502 | -34.407 | -19.3994
11.2563 | -34.4074 | -19.4088
11.2624 | -34.4079 | -19.418
11.2685 | -34.4083 | -19.4274
11.2746 | -34.4087 | -19.4369
11.2807 | -34.4092 | -19.4466
11.2868 | -34.4096 | -19.4563
11.2929 | -34.4101 | -19.4654
11.299 | -34.4105 | -19.4749
11.3051 | -34.411 | -19.4844
11.3112 | -34.4114 | -19.4939
11.3173 | -34.4119 | -19.5031
11.3234 | -34.4124 | -19.5123
11.3295 | -34.4128 | -19.5218
11.3356 | -34.4132 | -19.5313
11.3417 | -34.4137 | -19.5406
11.3478 | -34.4141 | -19.5501
11.3539 | -34.4145 | -19.5596
11.36 | -34.415 | -19.5688
11.3661 | -34.4155 | -19.5781
11.3722 | -34.4159 | -19.5873
11.3783 | -34.4164 | -19.5968
11.3844 | -34.4168 | -19.6062
11.3906 | -34.4173 | -19.6156
11.3967 | -34.4177 | -19.625
11.4028 | -34.4182 | -19.6345
11.4089 | -34.4186 | -19.6436
11.415 | -34.419 | -19.653
11.4211 | -34.4195 | -19.6624
11.4272 | -34.4199 | -19.6719
11.4333 | -34.4204 | -19.681
11.4394 | -34.4208 | -19.6902
11.4455 | -34.4212 | -19.6997
11.4516 | -34.4217 | -19.709
11.4577 | -34.4221 | -19.7184
11.4638 | -34.4226 | -19.7276
11.4699 | -34.4231 | -19.737
11.476 | -34.4235 | -19.7465
11.4821 | -34.4239 | -19.756
11.4882 | -34.4244 | -19.7653
11.4943 | -34.4249 | -19.7748
11.5004 | -34.4253 | -19.7839
11.5065 | -34.4258 | -19.7932
11.5126 | -34.4262 | -19.8027
11.5187 | -34.4267 | -19.8121
11.5248 | -34.4271 | -19.8219
11.5309 | -34.4276 | -19.8311
11.5371 | -34.428 | -19.8404
11.5432 | -34.4285 | -19.8499
11.5493 | -34.4289 | -19.8595
11.5554 | -34.4294 | -19.8688
11.5615 | -34.4298 | -19.8784
11.5676 | -34.4303 | -19.8878
11.5737 | -34.4307 | -19.8971
11.5798 | -34.4312 | -19.9063
11.5859 | -34.4317 | -19.9156
11.592 | -34.4321 | -19.9249
11.5981 | -34.4326 | -19.9343
11.6042 | -34.433 | -19.9436
11.6103 | -34.4335 | -19.9528
11.6164 | -34.4339 | -19.9621
11.6225 | -34.4344 | -19.9715
11.6286 | -34.4348 | -19.9808
11.6347 | -34.4353 | -19.9902
11.6408 | -34.4357 | -19.9995
11.6469 | -34.4362 | -20.009
11.653 | -34.4366 | -20.0182
11.6591 | -34.4371 | -20.0276
11.6652 | -34.4375 | -20.0368
11.6713 | -34.438 | -20.0461
11.6775 | -34.4385 | -20.0554
11.6836 | -34.439 | -20.0647
11.6897 | -34.4394 | -20.074
11.6958 | -34.4399 | -20.0835
11.7019 | -34.4403 | -20.0928
11.708 | -34.4408 | -20.1021
11.7141 | -34.4413 | -20.1112
11.7202 | -34.4417 | -20.1206
11.7263 | -34.4422 | -20.1299
11.7324 | -34.4426 | -20.1393
11.7385 | -34.4431 | -20.1487
11.7446 | -34.4435 | -20.1579
11.7507 | -34.444 | -20.1671
11.7568 | -34.4445 | -20.1764
11.7629 | -34.4449 | -20.1858
11.769 | -34.4454 | -20.1949
11.7751 | -34.4458 | -20.2042
11.7812 | -34.4463 | -20.2134
11.7873 | -34.4468 | -20.2227
11.7934 | -34.4472 | -20.2321
11.7995 | -34.4477 | -20.2413
11.8056 | -34.4482 | -20.2505
11.8117 | -34.4486 | -20.2598
11.8178 | -34.4491 | -20.2689
11.824 | -34.4496 | -20.2781
11.8301 | -34.45 | -20.2873
11.8362 | -34.4505 | -20.2964
11.8423 | -34.4509 | -20.3057
11.8484 | -34.4514 | -20.315
11.8545 | -34.4519 | -20.3241
11.8606 | -34.4524 | -20.3333
11.8667 | -34.4528 | -20.3425
11.8728 | -34.4532 | -20.3515
11.8789 | -34.4537 | -20.3607
11.885 | -34.4541 | -20.37
11.8911 | -34.4546 | -20.3792
11.8972 | -34.4551 | -20.3884
11.9033 | -34.4555 | -20.3977
11.9094 | -34.456 | -20.4067
11.9155 | -34.4564 | -20.4159
11.9216 | -34.4569 | -20.4249
11.9277 | -34.4574 | -20.4341
11.9338 | -34.4578 | -20.4433
11.9399 | -34.4583 | -20.4526
11.946 | -34.4588 | -20.4618
11.9521 | -34.4593 | -20.4709
11.9582 | -34.4598 | -20.48
11.9644 | -34.4602 | -20.489
11.9705 | -34.4608 | -20.4983
11.9766 | -34.4612 | -20.5076
11.9827 | -34.4617 | -20.5167
11.9888 | -34.4622 | -20.5258
11.9949 | -34.4627 | -20.5347
12.001 | -34.4632 | -20.544
12.0071 | -34.4636 | -20.5531
12.0132 | -34.4641 | -20.5623
12.0193 | -34.4646 | -20.5716
12.0254 | -34.465 | -20.5808
12.0315 | -34.4655 | -20.5897
12.0376 | -34.466 | -20.5987
12.0437 | -34.4664 | -20.6078
12.0498 | -34.4669 | -20.617
12.0559 | -34.4674 | -20.626
12.062 | -34.4679 | -20.6349
12.0681 | -34.4684 | -20.6441
12.0742 | -34.4688 | -20.6531
12.0803 | -34.4693 | -20.6624
12.0864 | -34.4698 | -20.6716
12.0925 | -34.4703 | -20.6807
12.0986 | -34.4708 | -20.6897
12.1047 | -34.4712 | -20.6989
12.1109 | -34.4717 | -20.7079
12.117 | -34.4722 | -20.717
12.1231 | -34.4727 | -20.7261
12.1292 | -34.4732 | -20.7352
12.1353 | -34.4737 | -20.7442
12.1414 | -34.4741 | -20.7533
12.1475 | -34.4746 | -20.7626
12.1536 | -34.4751 | -20.7718
12.1597 | -34.4755 | -20.7807
12.1658 | -34.476 | -20.7899
12.1719 | -34.4765 | -20.7987
12.178 | -34.477 | -20.8079
12.1841 | -34.4775 | -20.8171
12.1902 | -34.478 | -20.8264
12.1963 | -34.4785 | -20.8355
12.2024 | -34.479 | -20.8443
12.2085 | -34.4794 | -20.8535
12.2146 | -34.4799 | -20.8626
12.2207 | -34.4804 | -20.8717
12.2268 | -34.4809 | -20.8807
12.2329 | -34.4814 | -20.8896
12.239 | -34.4819 | -20.8983
12.2451 | -34.4824 | -20.9073
12.2513 | -34.4829 | -20.9165
12.2574 | -34.4833 | -20.9255
12.2635 | -34.4838 | -20.9346
12.2696 | -34.4843 | -20.9436
12.2757 | -34.4848 | -20.9524
12.2818 | -34.4853 | -20.9614
12.2879 | -34.4858 | -20.9705
12.294 | -34.4863 | -20.9794
12.3001 | -34.4868 | -20.9884
12.3062 | -34.4873 | -20.9974
12.3123 | -34.4878 | -21.006
12.3184 | -34.4883 | -21.0148
12.3245 | -34.4887 | -21.0237
12.3306 | -34.4892 | -21.0328
12.3367 | -34.4897 | -21.0418
12.3428 | -34.4902 | -21.0507
12.3489 | -34.4907 | -21.0594
12.355 | -34.4912 | -21.0685
12.3611 | -34.4917 | -21.0777
12.3672 | -34.4921 | -21.0866
12.3733 | -34.4926 | -21.0956
12.3794 | -34.4931 | -21.1042
12.3855 | -34.4936 | -21.1131
12.3916 | -34.4941 | -21.1221
12.3978 | -34.4946 | -21.1311
12.4039 | -34.4951 | -21.1402
12.41 | -34.4956 | -21.1491
12.4161 | -34.4961 | -21.1576
12.4222 | -34.4966 | -21.1666
12.4283 | -34.4971 | -21.1754
12.4344 | -34.4976 | -21.1845
12.4405 | -34.498 | -21.1935
12.4466 | -34.4985 | -21.2025
12.4527 | -34.499 | -21.2112
12.4588 | -34.4995 | -21.2203
12.4649 | -34.5 | -21.2293
12.471 | -34.5005 | -21.2385
12.4771 | -34.501 | -21.2473
12.4832 | -34.5015 | -21.2563
12.4893 | -34.502 | -21.2651
12.4954 | -34.5024 | -21.2738
12.5015 | -34.5029 | -21.2825
12.5076 | -34.5034 | -21.2914
12.5137 | -34.504 | -21.3002
12.5198 | -34.5045 | -21.309
12.5259 | -34.5049 | -21.3179
12.532 | -34.5054 | -21.3268
12.5382 | -34.5059 | -21.3357
12.5443 | -34.5064 | -21.3446
12.5504 | -34.5069 | -21.3535
12.5565 | -34.5074 | -21.3623
12.5626 | -34.5079 | -21.3713
12.5687 | -34.5084 | -21.3802
12.5748 | -34.5089 | -21.3891
12.5809 | -34.5094 | -21.3979
12.587 | -34.5099 | -21.4069
12.5931 | -34.5104 | -21.4154
12.5992 | -34.5109 | -21.4243
12.6053 | -34.5114 | -21.4332
12.6114 | -34.5119 | -21.4421
12.6175 | -34.5124 | -21.451
12.6236 | -34.5129 | -21.4599
12.6297 | -34.5134 | -21.4688
12.6358 | -34.5139 | -21.4774
12.6419 | -34.5144 | -21.4862
12.648 | -34.5149 | -21.4952
12.6541 | -34.5154 | -21.504
12.6602 | -34.5159 | -21.5123
12.6663 | -34.5164 | -21.521
12.6724 | -34.5169 | -21.5298
12.6785 | -34.5174 | -21.5387
12.6847 | -34.5179 | -21.5474
12.6908 | -34.5184 | -21.5561
12.6969 | -34.5189 | -21.5646
12.703 | -34.5194 | -21.5734
12.7091 | -34.5199 | -21.5823
12.7152 | -34.5203 | -21.5909
12.7213 | -34.5208 | -21.5998
12.7274 | -34.5213 | -21.6087
12.7335 | -34.5219 | -21.6171
12.7396 | -34.5223 | -21.6258
12.7457 | -34.5228 | -21.6348
12.7518 | -34.5233 | -21.6435
12.7579 | -34.5238 | -21.6523
12.764 | -34.5243 | -21.6612
12.7701 | -34.5248 | -21.6698
12.7762 | -34.5253 | -21.6785
12.7823 | -34.5258 | -21.6873
12.7884 | -34.5263 | -21.696
12.7945 | -34.5268 | -21.705
12.8006 | -34.5273 | -21.7137
12.8067 | -34.5279 | -21.7224
12.8128 | -34.5283 | -21.731
12.8189 | -34.5289 | -21.7397
12.8251 | -34.5294 | -21.7487
12.8312 | -34.5298 | -21.7575
12.8373 | -34.5303 | -21.7662
12.8434 | -34.5309 | -21.7747
12.8495 | -34.5313 | -21.7833
12.8556 | -34.5318 | -21.792
12.8617 | -34.5324 | -21.8007
12.8678 | -34.5329 | -21.8097
12.8739 | -34.5334 | -21.8182
12.88 | -34.5338 | -21.8266
12.8861 | -34.5343 | -21.8352
12.8922 | -34.5348 | -21.8439
12.8983 | -34.5353 | -21.8527
12.9044 | -34.5358 | -21.861
12.9105 | -34.5363 | -21.8696
12.9166 | -34.5368 | -21.8782
12.9227 | -34.5373 | -21.8868
12.9288 | -34.5378 | -21.8955
12.9349 | -34.5383 | -21.9043
12.941 | -34.5388 | -21.9129
12.9471 | -34.5394 | -21.9215
12.9532 | -34.5399 | -21.9303
12.9593 | -34.5404 | -21.939
12.9654 | -34.5409 | -21.9476
12.9716 | -34.5414 | -21.9565
12.9777 | -34.5419 | -21.9651
12.9838 | -34.5424 | -21.9736
12.9899 | -34.5429 | -21.9823
12.996 | -34.5434 | -21.9909
13.0021 | -34.5439 | -21.9996
13.0082 | -34.5445 | -22.008
13.0143 | -34.545 | -22.0164
13.0204 | -34.5454 | -22.0248
13.0265 | -34.546 | -22.0335
13.0326 | -34.5465 | -22.0422
13.0387 | -34.547 | -22.0509
13.0448 | -34.5474 | -22.0596
13.0509 | -34.548 | -22.068
13.057 | -34.5485 | -22.0765
13.0631 | -34.549 | -22.0851
13.0692 | -34.5495 | -22.0935
13.0753 | -34.5501 | -22.1023
13.0814 | -34.5505 | -22.1109
13.0875 | -34.5511 | -22.1191
13.0936 | -34.5516 | -22.1277
13.0997 | -34.5521 | -22.1363
13.1058 | -34.5526 | -22.1449
13.112 | -34.5531 | -22.1534
13.1181 | -34.5537 | -22.1619
13.1242 | -34.5541 | -22.1705
13.1303 | -34.5546 | -22.179
13.1364 | -34.5551 | -22.1877
13.1425 | -34.5556 | -22.1962
13.1486 | -34.5561 | -22.2049
13.1547 | -34.5566 | -22.2132
13.1608 | -34.5571 | -22.2216
13.1669 | -34.5576 | -22.23
13.173 | -34.5581 | -22.2387
13.1791 | -34.5587 | -22.2471
13.1852 | -34.5592 | -22.2556
13.1913 | -34.5597 | -22.2641
13.1974 | -34.5602 | -22.2727
13.2035 | -34.5607 | -22.2812
13.2096 | -34.5612 | -22.2897
13.2157 | -34.5617 | -22.2982
13.2218 | -34.5622 | -22.3067
13.2279 | -34.5627 | -22.315
13.234 | -34.5632 | -22.3235
13.2401 | -34.5637 | -22.3319
13.2462 | -34.5642 | -22.3405
13.2524 | -34.5647 | -22.3492
13.2585 | -34.5652 | -22.3578
13.2646 | -34.5657 | -22.3663
13.2707 | -34.5662 | -22.3749
13.2768 | -34.5668 | -22.3835
13.2829 | -34.5672 | -22.3923
13.289 | -34.5677 | -22.4008
13.2951 | -34.5683 | -22.4092
13.3012 | -34.5688 | -22.4177
13.3073 | -34.5693 | -22.4264
13.3134 | -34.5698 | -22.4348
13.3195 | -34.5703 | -22.4434
13.3256 | -34.5708 | -22.4517
13.3317 | -34.5713 | -22.4599
13.3378 | -34.5718 | -22.4684
13.3439 | -34.5723 | -22.4772
13.35 | -34.5728 | -22.4857
13.3561 | -34.5733 | -22.4939
13.3622 | -34.5739 | -22.5023
13.3683 | -34.5743 | -22.5104
13.3744 | -34.5748 | -22.5186
13.3805 | -34.5753 | -22.527
13.3866 | -34.5758 | -22.5355
13.3927 | -34.5763 | -22.5441
13.3989 | -34.5768 | -22.5523
13.405 | -34.5773 | -22.5606
13.4111 | -34.5778 | -22.5689
13.4172 | -34.5783 | -22.5772
13.4233 | -34.5788 | -22.5858
13.4294 | -34.5793 | -22.5946
13.4355 | -34.5798 | -22.6029
13.4416 | -34.5803 | -22.6113
13.4477 | -34.5809 | -22.6197
13.4538 | -34.5814 | -22.6282
13.4599 | -34.5819 | -22.6367
13.466 | -34.5824 | -22.6452
13.4721 | -34.5829 | -22.6536
13.4782 | -34.5834 | -22.6621
13.4843 | -34.5839 | -22.6708
13.4904 | -34.5844 | -22.6791
13.4965 | -34.5849 | -22.6875
13.5026 | -34.5854 | -22.6961
13.5087 | -34.5859 | -22.7042
13.5148 | -34.5865 | -22.7129
13.5209 | -34.5869 | -22.7215
13.527 | -34.5875 | -22.73
13.5331 | -34.588 | -22.7383
13.5393 | -34.5885 | -22.7463
13.5454 | -34.589 | -22.7548
13.5515 | -34.5896 | -22.7635
13.5576 | -34.5901 | -22.7718
13.5637 | -34.5906 | -22.7803
13.5698 | -34.5911 | -22.7888
13.5759 | -34.5916 | -22.7968
13.582 | -34.5921 | -22.8051
13.5881 | -34.5926 | -22.8136
13.5942 | -34.5931 | -22.8223
13.6003 | -34.5936 | -22.8307
13.6064 | -34.5941 | -22.8394
13.6125 | -34.5946 | -22.8479
13.6186 | -34.5952 | -22.8562
13.6247 | -34.5957 | -22.8647
13.6308 | -34.5961 | -22.8731
13.6369 | -34.5966 | -22.8818
13.643 | -34.5971 | -22.8904
13.6491 | -34.5976 | -22.8989
13.6552 | -34.5981 | -22.9073
13.6613 | -34.5986 | -22.9157
13.6674 | -34.5991 | -22.9241
13.6735 | -34.5996 | -22.9326
13.6796 | -34.6001 | -22.9409
13.6858 | -34.6006 | -22.9492
13.6919 | -34.6011 | -22.9576
13.698 | -34.6016 | -22.966
13.7041 | -34.6021 | -22.9745
13.7102 | -34.6027 | -22.9831
13.7163 | -34.6032 | -22.9914
13.7224 | -34.6037 | -22.9998
13.7285 | -34.6042 | -23.0082
13.7346 | -34.6047 | -23.0166
13.7407 | -34.6052 | -23.0251
13.7468 | -34.6057 | -23.0335
13.7529 | -34.6062 | -23.0418
13.759 | -34.6067 | -23.0502
13.7651 | -34.6072 | -23.0587
13.7712 | -34.6077 | -23.0671
13.7773 | -34.6083 | -23.0755
13.7834 | -34.6087 | -23.0837
13.7895 | -34.6093 | -23.0921
13.7956 | -34.6098 | -23.1006
13.8017 | -34.6103 | -23.1089
13.8078 | -34.6108 | -23.1174
13.8139 | -34.6113 | -23.1257
13.82 | -34.6118 | -23.1339
13.8262 | -34.6122 | -23.1424
13.8323 | -34.6128 | -23.151
13.8384 | -34.6133 | -23.1596
13.8445 | -34.6138 | -23.1678
13.8506 | -34.6143 | -23.1762
13.8567 | -34.6148 | -23.1846
13.8628 | -34.6153 | -23.193
13.8689 | -34.6158 | -23.2016
13.875 | -34.6163 | -23.2099
13.8811 | -34.6168 | -23.2183
13.8872 | -34.6172 | -23.2268
13.8933 | -34.6178 | -23.2351
13.8994 | -34.6183 | -23.2435
13.9055 | -34.6188 | -23.2518
13.9116 | -34.6193 | -23.2602
13.9177 | -34.6197 | -23.2686
13.9238 | -34.6202 | -23.2769
13.9299 | -34.6208 | -23.2852
13.936 | -34.6213 | -23.2937
13.9421 | -34.6218 | -23.302
13.9482 | -34.6223 | -23.3101
13.9543 | -34.6228 | -23.3185
13.9604 | -34.6233 | -23.3269
13.9665 | -34.6238 | -23.3353
13.9727 | -34.6243 | -23.344
13.9788 | -34.6249 | -23.3525
13.9849 | -34.6254 | -23.361
13.991 | -34.6259 | -23.3696
13.9971 | -34.6264 | -23.3778
14.0032 | -34.6268 | -23.3862
14.0093 | -34.6274 | -23.3947
14.0154 | -34.6279 | -23.4029
14.0215 | -34.6284 | -23.4112
14.0276 | -34.6289 | -23.4195
14.0337 | -34.6294 | -23.4278
14.0398 | -34.6299 | -23.4361
14.0459 | -34.6305 | -23.4443
14.052 | -34.631 | -23.4528
14.0581 | -34.6314 | -23.4612
14.0642 | -34.6319 | -23.4696
14.0703 | -34.6324 | -23.478
14.0764 | -34.633 | -23.4861
14.0825 | -34.6335 | -23.4944
14.0886 | -34.634 | -23.5028
14.0947 | -34.6345 | -23.5112
14.1008 | -34.635 | -23.5196
14.1069 | -34.6355 | -23.5279
14.1131 | -34.636 | -23.5364
14.1192 | -34.6365 | -23.5446
14.1253 | -34.637 | -23.5528
14.1314 | -34.6375 | -23.5611
14.1375 | -34.638 | -23.5694
14.1436 | -34.6385 | -23.5778
14.1497 | -34.639 | -23.5861
14.1558 | -34.6395 | -23.5946
14.1619 | -34.64 | -23.6029
14.168 | -34.6405 | -23.6113
14.1741 | -34.6411 | -23.6197
14.1802 | -34.6416 | -23.6281
14.1863 | -34.6421 | -23.6365
14.1924 | -34.6426 | -23.6448
14.1985 | -34.6431 | -23.6532
14.2046 | -34.6436 | -23.6617
14.2107 | -34.6442 | -23.6699
14.2168 | -34.6447 | -23.6782
14.2229 | -34.6451 | -23.6866
14.229 | -34.6456 | -23.6951
14.2351 | -34.6461 | -23.7035
14.2412 | -34.6466 | -23.7119
14.2473 | -34.6472 | -23.7203
14.2534 | -34.6477 | -23.7284
14.2596 | -34.6482 | -23.7366
14.2657 | -34.6487 | -23.7451
14.2718 | -34.6492 | -23.7533
14.2779 | -34.6497 | -23.7617
14.284 | -34.6503 | -23.7701
14.2901 | -34.6508 | -23.7784
14.2962 | -34.6513 | -23.7868
14.3023 | -34.6518 | -23.7952
14.3084 | -34.6523 | -23.8036
14.3145 | -34.6528 | -23.8121
14.3206 | -34.6533 | -23.8208
14.3267 | -34.6538 | -23.8291
14.3328 | -34.6544 | -23.8374
14.3389 | -34.6549 | -23.8456
14.345 | -34.6554 | -23.8538
14.3511 | -34.6559 | -23.8624
14.3572 | -34.6564 | -23.8708
14.3633 | -34.6569 | -23.8792
14.3694 | -34.6575 | -23.8875
14.3755 | -34.658 | -23.8956
14.3816 | -34.6585 | -23.9041
14.3877 | -34.659 | -23.9123
14.3938 | -34.6595 | -23.9208
14.4 | -34.6601 | -23.9292
14.4061 | -34.6606 | -23.9375
14.4122 | -34.6611 | -23.946
14.4183 | -34.6616 | -23.9544
14.4244 | -34.6621 | -23.9628
14.4305 | -34.6627 | -23.971
14.4366 | -34.6632 | -23.9793
14.4427 | -34.6637 | -23.9879
14.4488 | -34.6642 | -23.9964
14.4549 | -34.6647 | -24.0046
14.461 | -34.6652 | -24.0128
14.4671 | -34.6657 | -24.0211
14.4732 | -34.6663 | -24.0296
14.4793 | -34.6668 | -24.038
14.4854 | -34.6673 | -24.0463
14.4915 | -34.6678 | -24.0546
14.4976 | -34.6683 | -24.0627
14.5037 | -34.6689 | -24.0708
14.5098 | -34.6694 | -24.079
14.5159 | -34.6699 | -24.0874
14.522 | -34.6704 | -24.0958
14.5281 | -34.6709 | -24.1041
14.5342 | -34.6714 | -24.1125
14.5403 | -34.6719 | -24.1208
14.5465 | -34.6724 | -24.1292
14.5526 | -34.6729 | -24.1377
14.5587 | -34.6735 | -24.1461
14.5648 | -34.674 | -24.1543
14.5709 | -34.6745 | -24.1627
14.577 | -34.675 | -24.1709
14.5831 | -34.6755 | -24.1794
14.5892 | -34.676 | -24.188
14.5953 | -34.6765 | -24.1962
14.6014 | -34.677 | -24.2047
14.6075 | -34.6775 | -24.213
14.6136 | -34.678 | -24.2212
14.6197 | -34.6785 | -24.2296
14.6258 | -34.679 | -24.238
14.6319 | -34.6795 | -24.2463
14.638 | -34.68 | -24.2545
14.6441 | -34.6805 | -24.2629
14.6502 | -34.681 | -24.2713
14.6563 | -34.6816 | -24.2797
14.6624 | -34.6821 | -24.288
14.6685 | -34.6826 | -24.2963
14.6746 | -34.6831 | -24.3044
14.6807 | -34.6836 | -24.3126
14.6869 | -34.6841 | -24.321
14.693 | -34.6847 | -24.3293
14.6991 | -34.6851 | -24.3376
14.7052 | -34.6857 | -24.3456
14.7113 | -34.6862 | -24.354
14.7174 | -34.6867 | -24.3624
14.7235 | -34.6872 | -24.3708
14.7296 | -34.6877 | -24.3791
14.7357 | -34.6882 | -24.3874
14.7418 | -34.6888 | -24.3956
14.7479 | -34.6893 | -24.4039
14.754 | -34.6898 | -24.4124
14.7601 | -34.6903 | -24.4207
14.7662 | -34.6909 | -24.4289
14.7723 | -34.6914 | -24.4371
14.7784 | -34.6919 | -24.4455
14.7845 | -34.6924 | -24.4538
14.7906 | -34.6929 | -24.4623
14.7967 | -34.6934 | -24.4706
14.8028 | -34.6939 | -24.4787
14.8089 | -34.6945 | -24.487
14.815 | -34.695 | -24.4953
14.8211 | -34.6955 | -24.5037
14.8272 | -34.696 | -24.5121
14.8334 | -34.6965 | -24.5204
14.8395 | -34.6971 | -24.5287
14.8456 | -34.6976 | -24.5369
14.8517 | -34.6981 | -24.5453
14.8578 | -34.6986 | -24.5537
14.8639 | -34.6991 | -24.5619
14.87 | -34.6997 | -24.5702
14.8761 | -34.7002 | -24.5784
14.8822 | -34.7007 | -24.5868
14.8883 | -34.7012 | -24.5952
14.8944 | -34.7016 | -24.6036
14.9005 | -34.7022 | -24.6119
14.9066 | -34.7027 | -24.6205
14.9127 | -34.7032 | -24.6288
14.9188 | -34.7037 | -24.6372
14.9249 | -34.7042 | -24.6456
14.931 | -34.7048 | -24.654
14.9371 | -34.7053 | -24.6622
14.9432 | -34.7058 | -24.6705
14.9493 | -34.7063 | -24.6786
14.9554 | -34.7068 | -24.687
14.9615 | -34.7073 | -24.6954
14.9676 | -34.7078 | -24.7039
14.9738 | -34.7083 | -24.7121
14.9799 | -34.7088 | -24.7205
14.986 | -34.7094 | -24.7287
14.9921 | -34.7099 | -24.7369
14.9982 | -34.7104 | -24.745
15.0043 | -34.7109 | -24.7533
15.0104 | -34.7114 | -24.7617
15.0165 | -34.7119 | -24.7702
15.0226 | -34.7125 | -24.7783
15.0287 | -34.713 | -24.7867
15.0348 | -34.7135 | -24.795
15.0409 | -34.714 | -24.8034
15.047 | -34.7145 | -24.8118
15.0531 | -34.7151 | -24.8201
15.0592 | -34.7156 | -24.8284
15.0653 | -34.7162 | -24.8365
15.0714 | -34.7167 | -24.8447
15.0775 | -34.7171 | -24.8528
15.0836 | -34.7177 | -24.861
15.0897 | -34.7182 | -24.8694
15.0958 | -34.7188 | -24.8775
15.1019 | -34.7193 | -24.8859
15.108 | -34.7198 | -24.8941
15.1141 | -34.7203 | -24.9024
15.1203 | -34.7208 | -24.9109
15.1264 | -34.7213 | -24.919
15.1325 | -34.7218 | -24.9272
15.1386 | -34.7224 | -24.9355
15.1447 | -34.7229 | -24.9436
15.1508 | -34.7234 | -24.9519
15.1569 | -34.724 | -24.9603
15.163 | -34.7245 | -24.9685
15.1691 | -34.725 | -24.9769
15.1752 | -34.7255 | -24.9852
15.1813 | -34.726 | -24.9935
15.1874 | -34.7266 | -25.0018
15.1935 | -34.7271 | -25.0099
15.1996 | -34.7276 | -25.0181
15.2057 | -34.7281 | -25.0263
15.2118 | -34.7286 | -25.0347
15.2179 | -34.7291 | -25.0432
15.224 | -34.7296 | -25.0514
15.2301 | -34.7302 | -25.0597
15.2362 | -34.7307 | -25.068
15.2423 | -34.7312 | -25.0765
15.2484 | -34.7317 | -25.0848
15.2545 | -34.7323 | -25.093
15.2607 | -34.7328 | -25.1014
15.2668 | -34.7333 | -25.1095
15.2729 | -34.7338 | -25.1177
15.279 | -34.7343 | -25.1262
15.2851 | -34.7349 | -25.1345
15.2912 | -34.7354 | -25.1429
15.2973 | -34.7359 | -25.1514
15.3034 | -34.7364 | -25.1596
15.3095 | -34.7369 | -25.1678
15.3156 | -34.7375 | -25.1761
15.3217 | -34.738 | -25.1844
15.3278 | -34.7386 | -25.1928
15.3339 | -34.7391 | -25.2009
15.34 | -34.7396 | -25.209
15.3461 | -34.7402 | -25.217
15.3522 | -34.7407 | -25.2255
15.3583 | -34.7412 | -25.2339
15.3644 | -34.7418 | -25.2419
15.3705 | -34.7423 | -25.25
15.3766 | -34.7428 | -25.2585
15.3827 | -34.7434 | -25.2665
15.3888 | -34.7439 | -25.2748
15.3949 | -34.7444 | -25.2832
15.401 | -34.745 | -25.2917
15.4072 | -34.7455 | -25.2998
15.4133 | -34.746 | -25.3083
15.4194 | -34.7465 | -25.3165
15.4255 | -34.7471 | -25.3248
15.4316 | -34.7476 | -25.3331
15.4377 | -34.7481 | -25.3415
15.4438 | -34.7487 | -25.3498
15.4499 | -34.7492 | -25.3579
15.456 | -34.7498 | -25.3663
15.4621 | -34.7503 | -25.3746
15.4682 | -34.7508 | -25.383
15.4743 | -34.7514 | -25.3911
15.4804 | -34.7519 | -25.3992
15.4865 | -34.7524 | -25.4075
15.4926 | -34.753 | -25.4159
15.4987 | -34.7535 | -25.4241
15.5048 | -34.754 | -25.4322
15.5109 | -34.7546 | -25.4405
15.517 | -34.7552 | -25.4487
15.5231 | -34.7557 | -25.457
15.5292 | -34.7562 | -25.4652
15.5353 | -34.7567 | -25.4737
15.5414 | -34.7572 | -25.4819
15.5476 | -34.7578 | -25.49
15.5537 | -34.7583 | -25.4982
15.5598 | -34.7588 | -25.5067
15.5659 | -34.7594 | -25.5149
15.572 | -34.7599 | -25.523
15.5781 | -34.7604 | -25.5312
15.5842 | -34.761 | -25.5395
15.5903 | -34.7615 | -25.5476
15.5964 | -34.7621 | -25.5557
15.6025 | -34.7626 | -25.5637
15.6086 | -34.7631 | -25.5719
15.6147 | -34.7637 | -25.5802
15.6208 | -34.7642 | -25.5882
15.6269 | -34.7648 | -25.5964
15.633 | -34.7653 | -25.6046
15.6391 | -34.7658 | -25.6129
15.6452 | -34.7664 | -25.6212
15.6513 | -34.7669 | -25.6292
15.6574 | -34.7674 | -25.6375
15.6635 | -34.768 | -25.6455
15.6696 | -34.7685 | -25.6538
15.6757 | -34.7691 | -25.6619
15.6818 | -34.7696 | -25.6703
15.688 | -34.7701 | -25.6785
15.6941 | -34.7707 | -25.6866
15.7002 | -34.7712 | -25.6949
15.7063 | -34.7717 | -25.703
15.7124 | -34.7722 | -25.7114
15.7185 | -34.7728 | -25.7196
15.7246 | -34.7733 | -25.7275
15.7307 | -34.7739 | -25.7356
15.7368 | -34.7744 | -25.7439
15.7429 | -34.7749 | -25.7519
15.749 | -34.7755 | -25.7602
15.7551 | -34.776 | -25.7685
15.7612 | -34.7766 | -25.7766
15.7673 | -34.7771 | -25.7849
15.7734 | -34.7776 | -25.793
15.7795 | -34.7782 | -25.8014
15.7856 | -34.7787 | -25.8096
15.7917 | -34.7793 | -25.8175
15.7978 | -34.7798 | -25.8259
15.8039 | -34.7803 | -25.8342
15.81 | -34.7808 | -25.8424
15.8161 | -34.7814 | -25.8505
15.8222 | -34.7819 | -25.8586
15.8283 | -34.7825 | -25.8666
15.8345 | -34.783 | -25.8749
15.8406 | -34.7836 | -25.8831
15.8467 | -34.7841 | -25.8913
15.8528 | -34.7846 | -25.8995
15.8589 | -34.7852 | -25.9077
15.865 | -34.7858 | -25.9157
15.8711 | -34.7863 | -25.9239
15.8772 | -34.7868 | -25.9321
15.8833 | -34.7874 | -25.9402
15.8894 | -34.7879 | -25.9486
15.8955 | -34.7884 | -25.9568
15.9016 | -34.789 | -25.965
15.9077 | -34.7895 | -25.9731
15.9138 | -34.79 | -25.9813
15.9199 | -34.7905 | -25.9894
15.926 | -34.7911 | -25.9974
15.9321 | -34.7916 | -26.0054
15.9382 | -34.7922 | -26.0134
15.9443 | -34.7927 | -26.0214
15.9504 | -34.7933 | -26.0297
15.9565 | -34.7938 | -26.0379
15.9626 | -34.7943 | -26.046
15.9687 | -34.7949 | -26.0538
15.9749 | -34.7955 | -26.0618
15.981 | -34.796 | -26.07
15.9871 | -34.7965 | -26.0783
15.9932 | -34.797 | -26.0862
15.9993 | -34.7976 | -26.0943
16.0054 | -34.7981 | -26.1023
16.0115 | -34.7987 | -26.1106
16.0176 | -34.7992 | -26.1187
16.0237 | -34.7998 | -26.1269
16.0298 | -34.8003 | -26.1349
16.0359 | -34.8008 | -26.1429
16.042 | -34.8014 | -26.1511
16.0481 | -34.8019 | -26.1592
16.0542 | -34.8025 | -26.1675
16.0603 | -34.803 | -26.1755
16.0664 | -34.8035 | -26.1837
16.0725 | -34.8041 | -26.1916
16.0786 | -34.8046 | -26.1996
16.0847 | -34.8052 | -26.2079
16.0908 | -34.8057 | -26.216
16.0969 | -34.8062 | -26.2244
16.103 | -34.8067 | -26.2325
16.1091 | -34.8073 | -26.2405
16.1152 | -34.8079 | -26.2485
16.1214 | -34.8084 | -26.2567
16.1275 | -34.8089 | -26.2648
16.1336 | -34.8095 | -26.2729
16.1397 | -34.81 | -26.2809
16.1458 | -34.8106 | -26.2889
16.1519 | -34.8111 | -26.2971
16.158 | -34.8116 | -26.3053
16.1641 | -34.8122 | -26.3135
16.1702 | -34.8127 | -26.3216
16.1763 | -34.8133 | -26.3299
16.1824 | -34.8138 | -26.3379
16.1885 | -34.8143 | -26.346
16.1946 | -34.8149 | -26.3539
16.2007 | -34.8155 | -26.3621
16.2068 | -34.816 | -26.37
16.2129 | -34.8166 | -26.3784
16.219 | -34.8171 | -26.3865
16.2251 | -34.8177 | -26.3944
16.2312 | -34.8182 | -26.4023
16.2373 | -34.8187 | -26.4103
16.2434 | -34.8193 | -26.4184
16.2495 | -34.8198 | -26.426
16.2556 | -34.8204 | -26.4344
16.2618 | -34.8209 | -26.4426
16.2679 | -34.8214 | -26.4507
16.274 | -34.822 | -26.459
16.2801 | -34.8225 | -26.4669
16.2862 | -34.8231 | -26.4748
16.2923 | -34.8236 | -26.4832
16.2984 | -34.8241 | -26.4911
16.3045 | -34.8246 | -26.4992
16.3106 | -34.8252 | -26.5073
16.3167 | -34.8258 | -26.5154
16.3228 | -34.8263 | -26.5232
16.3289 | -34.8268 | -26.5313
16.335 | -34.8273 | -26.5394
16.3411 | -34.8279 | -26.5473
16.3472 | -34.8285 | -26.5555
16.3533 | -34.829 | -26.5634
16.3594 | -34.8295 | -26.5711
16.3655 | -34.83 | -26.5791
16.3716 | -34.8306 | -26.587
16.3777 | -34.8312 | -26.5949
16.3838 | -34.8317 | -26.6029
16.3899 | -34.8323 | -26.6107
16.396 | -34.8328 | -26.6187
16.4021 | -34.8333 | -26.6268
16.4083 | -34.8339 | -26.6348
16.4144 | -34.8344 | -26.6429
16.4205 | -34.835 | -26.651
16.4266 | -34.8355 | -26.6591
16.4327 | -34.836 | -26.6672
16.4388 | -34.8366 | -26.6753
16.4449 | -34.8372 | -26.6832
16.451 | -34.8378 | -26.691
16.4571 | -34.8383 | -26.699
16.4632 | -34.8388 | -26.7066
16.4693 | -34.8394 | -26.7146
16.4754 | -34.8399 | -26.7229
16.4815 | -34.8405 | -26.7309
16.4876 | -34.841 | -26.7388
16.4937 | -34.8416 | -26.7467
16.4998 | -34.8421 | -26.7542
16.5059 | -34.8427 | -26.7621
16.512 | -34.8432 | -26.7703
16.5181 | -34.8438 | -26.7784
16.5242 | -34.8443 | -26.7866
16.5303 | -34.8449 | -26.7942
16.5364 | -34.8454 | -26.8024
16.5425 | -34.8459 | -26.8104
16.5487 | -34.8464 | -26.8184
16.5548 | -34.847 | -26.8263
16.5609 | -34.8476 | -26.8342
16.567 | -34.8481 | -26.8419
16.5731 | -34.8487 | -26.85
16.5792 | -34.8492 | -26.8578
16.5853 | -34.8498 | -26.866
16.5914 | -34.8503 | -26.874
16.5975 | -34.8509 | -26.8819
16.6036 | -34.8514 | -26.8897
16.6097 | -34.8519 | -26.8978
16.6158 | -34.8525 | -26.9058
16.6219 | -34.8531 | -26.9137
16.628 | -34.8536 | -26.9219
16.6341 | -34.8541 | -26.9297
16.6402 | -34.8547 | -26.9376
16.6463 | -34.8552 | -26.9458
16.6524 | -34.8558 | -26.9537
16.6585 | -34.8564 | -26.9618
16.6646 | -34.8569 | -26.9698
16.6707 | -34.8575 | -26.9778
16.6768 | -34.858 | -26.9857
16.6829 | -34.8585 | -26.9935
16.689 | -34.8591 | -27.0016
16.6952 | -34.8596 | -27.0095
16.7013 | -34.8602 | -27.0176
16.7074 | -34.8607 | -27.0254
16.7135 | -34.8613 | -27.0334
16.7196 | -34.8618 | -27.0413
16.7257 | -34.8624 | -27.0491
16.7318 | -34.8629 | -27.0569
16.7379 | -34.8635 | -27.0649
16.744 | -34.864 | -27.0725
16.7501 | -34.8646 | -27.0804
16.7562 | -34.8651 | -27.0882
16.7623 | -34.8656 | -27.096
16.7684 | -34.8662 | -27.104
16.7745 | -34.8668 | -27.1121
16.7806 | -34.8673 | -27.1197
16.7867 | -34.8679 | -27.1273
16.7928 | -34.8684 | -27.1355
16.7989 | -34.869 | -27.1435
16.805 | -34.8696 | -27.1515
16.8111 | -34.8701 | -27.1591
16.8172 | -34.8707 | -27.167
16.8233 | -34.8713 | -27.1751
16.8294 | -34.8718 | -27.183
16.8356 | -34.8724 | -27.1907
16.8417 | -34.8729 | -27.1984
16.8478 | -34.8735 | -27.206
16.8539 | -34.8741 | -27.214
16.86 | -34.8746 | -27.222
16.8661 | -34.8752 | -27.2299
16.8722 | -34.8757 | -27.2379
16.8783 | -34.8763 | -27.2457
16.8844 | -34.8769 | -27.2534
16.8905 | -34.8774 | -27.2613
16.8966 | -34.878 | -27.2691
16.9027 | -34.8785 | -27.277
16.9088 | -34.8791 | -27.2849
16.9149 | -34.8796 | -27.2928
16.921 | -34.8802 | -27.3005
16.9271 | -34.8807 | -27.3087
16.9332 | -34.8813 | -27.3166
16.9393 | -34.8818 | -27.3245
16.9454 | -34.8824 | -27.3322
16.9515 | -34.8829 | -27.3401
16.9576 | -34.8835 | -27.3476
16.9637 | -34.884 | -27.3555
16.9698 | -34.8846 | -27.3632
16.9759 | -34.8851 | -27.3709
16.9821 | -34.8856 | -27.3789
16.9882 | -34.8862 | -27.3864
16.9943 | -34.8867 | -27.3942
17.0004 | -34.8873 | -27.4021
17.0065 | -34.8879 | -27.41
17.0126 | -34.8884 | -27.4177
17.0187 | -34.8889 | -27.4257
17.0248 | -34.8895 | -27.4334
17.0309 | -34.89 | -27.4414
17.037 | -34.8906 | -27.4491
17.0431 | -34.8911 | -27.4568
17.0492 | -34.8917 | -27.4648
17.0553 | -34.8922 | -27.4725
17.0614 | -34.8928 | -27.4802
17.0675 | -34.8933 | -27.4879
17.0736 | -34.8939 | -27.4956
17.0797 | -34.8945 | -27.5032
17.0858 | -34.895 | -27.5112
17.0919 | -34.8956 | -27.5189
17.098 | -34.8962 | -27.5264
17.1041 | -34.8967 | -27.5343
17.1102 | -34.8973 | -27.5422
17.1163 | -34.8978 | -27.55
17.1225 | -34.8984 | -27.5576
17.1286 | -34.8989 | -27.565
17.1347 | -34.8995 | -27.5729
17.1408 | -34.9001 | -27.5808
17.1469 | -34.9006 | -27.5886
17.153 | -34.9011 | -27.5964
17.1591 | -34.9017 | -27.604
17.1652 | -34.9023 | -27.6114
17.1713 | -34.9028 | -27.6193
17.1774 | -34.9034 | -27.6274
17.1835 | -34.9039 | -27.6353
17.1896 | -34.9044 | -27.6431
17.1957 | -34.905 | -27.651
17.2018 | -34.9056 | -27.6585
17.2079 | -34.9061 | -27.6663
17.214 | -34.9067 | -27.6743
17.2201 | -34.9072 | -27.6821
17.2262 | -34.9077 | -27.69
17.2323 | -34.9083 | -27.6978
17.2384 | -34.9089 | -27.7055
17.2445 | -34.9094 | -27.7131
17.2506 | -34.91 | -27.7208
17.2567 | -34.9105 | -27.7286
17.2628 | -34.9111 | -27.7365
17.269 | -34.9116 | -27.7442
17.2751 | -34.9122 | -27.752
17.2812 | -34.9127 | -27.7598
17.2873 | -34.9133 | -27.7676
17.2934 | -34.9138 | -27.7755
17.2995 | -34.9144 | -27.7834
17.3056 | -34.9149 | -27.7907
17.3117 | -34.9155 | -27.7987
17.3178 | -34.916 | -27.8063
17.3239 | -34.9166 | -27.8141
17.33 | -34.9171 | -27.8218
17.3361 | -34.9177 | -27.8295
17.3422 | -34.9182 | -27.837
17.3483 | -34.9187 | -27.8447
17.3544 | -34.9193 | -27.8526
17.3605 | -34.9199 | -27.8603
17.3666 | -34.9204 | -27.8679
17.3727 | -34.921 | -27.8757
17.3788 | -34.9215 | -27.8832
17.3849 | -34.9221 | -27.8909
17.391 | -34.9227 | -27.8986
17.3971 | -34.9232 | -27.9062
17.4032 | -34.9238 | -27.914
17.4094 | -34.9243 | -27.9214
17.4155 | -34.9249 | -27.929
17.4216 | -34.9254 | -27.9367
17.4277 | -34.9259 | -27.9446
17.4338 | -34.9265 | -27.9523
17.4399 | -34.9271 | -27.96
17.446 | -34.9277 | -27.9674
17.4521 | -34.9282 | -27.975
17.4582 | -34.9288 | -27.9823
17.4643 | -34.9293 | -27.9899
17.4704 | -34.9299 | -27.9977
17.4765 | -34.9304 | -28.0056
17.4826 | -34.931 | -28.0129
17.4887 | -34.9315 | -28.0207
17.4948 | -34.9321 | -28.0285
17.5009 | -34.9326 | -28.0362
17.507 | -34.9332 | -28.0439
17.5131 | -34.9338 | -28.0516
17.5192 | -34.9343 | -28.0591
17.5253 | -34.9349 | -28.0667
17.5314 | -34.9354 | -28.0745
17.5375 | -34.9359 | -28.0822
17.5436 | -34.9365 | -28.0901
17.5497 | -34.937 | -28.0976
17.5559 | -34.9376 | -28.105
17.562 | -34.9381 | -28.1125
17.5681 | -34.9387 | -28.1202
17.5742 | -34.9393 | -28.1277
17.5803 | -34.9398 | -28.1352
17.5864 | -34.9404 | -28.1427
17.5925 | -34.941 | -28.1504
17.5986 | -34.9415 | -28.158
17.6047 | -34.9421 | -28.1658
17.6108 | -34.9427 | -28.1733
17.6169 | -34.9432 | -28.1808
17.623 | -34.9438 | -28.1883
17.6291 | -34.9444 | -28.1959
17.6352 | -34.9449 | -28.2036
17.6413 | -34.9455 | -28.2113
17.6474 | -34.946 | -28.2189
17.6535 | -34.9466 | -28.2266
17.6596 | -34.9471 | -28.234
17.6657 | -34.9477 | -28.2415
17.6718 | -34.9482 | -28.2492
17.6779 | -34.9487 | -28.2566
17.684 | -34.9493 | -28.2642
17.6901 | -34.9498 | -28.272
17.6963 | -34.9504 | -28.2793
17.7024 | -34.9509 | -28.2868
17.7085 | -34.9515 | -28.2943
17.7146 | -34.952 | -28.3017
17.7207 | -34.9526 | -28.3092
17.7268 | -34.9532 | -28.3166
17.7329 | -34.9537 | -28.3242
17.739 | -34.9543 | -28.3318
17.7451 | -34.9549 | -28.3392
17.7512 | -34.9554 | -28.3468
17.7573 | -34.956 | -28.3545
17.7634 | -34.9566 | -28.3617
17.7695 | -34.9571 | -28.3694
17.7756 | -34.9577 | -28.377
17.7817 | -34.9582 | -28.3848
17.7878 | -34.9588 | -28.3924
17.7939 | -34.9593 | -28.3999
17.8 | -34.9599 | -28.4073
17.8061 | -34.9604 | -28.4147
17.8122 | -34.961 | -28.4222
17.8183 | -34.9616 | -28.4297
17.8244 | -34.9621 | -28.4373
17.8305 | -34.9627 | -28.445
17.8366 | -34.9632 | -28.4522
17.8428 | -34.9638 | -28.4599
17.8489 | -34.9644 | -28.4672
17.855 | -34.9649 | -28.4748
17.8611 | -34.9655 | -28.4822
17.8672 | -34.9661 | -28.4895
17.8733 | -34.9666 | -28.497
17.8794 | -34.9672 | -28.5045
17.8855 | -34.9677 | -28.5121
17.8916 | -34.9683 | -28.5197
17.8977 | -34.9688 | -28.5273
17.9038 | -34.9694 | -28.5344
17.9099 | -34.9699 | -28.5418
17.916 | -34.9705 | -28.5494
17.9221 | -34.9711 | -28.5569
17.9282 | -34.9716 | -28.5645
17.9343 | -34.9722 | -28.5719
17.9404 | -34.9727 | -28.579
17.9465 | -34.9733 | -28.5863
17.9526 | -34.9738 | -28.5939
17.9587 | -34.9743 | -28.6016
17.9648 | -34.9749 | -28.609
17.9709 | -34.9754 | -28.6164
17.977 | -34.976 | -28.6238
17.9832 | -34.9765 | -28.6314
17.9893 | -34.9771 | -28.6391
17.9954 | -34.9777 | -28.6467
18.0015 | -34.9782 | -28.6543
18.0076 | -34.9788 | -28.6616
18.0137 | -34.9793 | -28.6692
18.0198 | -34.9798 | -28.6767
18.0259 | -34.9803 | -28.6843
18.032 | -34.9809 | -28.6916
18.0381 | -34.9814 | -28.6993
18.0442 | -34.982 | -28.7067
18.0503 | -34.9826 | -28.7141
18.0564 | -34.9831 | -28.7217
18.0625 | -34.9837 | -28.7293
18.0686 | -34.9843 | -28.7367
18.0747 | -34.9848 | -28.744
18.0808 | -34.9854 | -28.7515
18.0869 | -34.9859 | -28.7589
18.093 | -34.9865 | -28.7665
18.0991 | -34.987 | -28.7738
18.1052 | -34.9875 | -28.7815
18.1113 | -34.9881 | -28.7889
18.1174 | -34.9886 | -28.7962
18.1236 | -34.9892 | -28.8038
18.1297 | -34.9898 | -28.8116
18.1358 | -34.9903 | -28.8189
18.1419 | -34.9908 | -28.8265
18.148 | -34.9914 | -28.8338
18.1541 | -34.9919 | -28.841
18.1602 | -34.9925 | -28.8486
18.1663 | -34.993 | -28.8559
18.1724 | -34.9936 | -28.8636
18.1785 | -34.9942 | -28.871
18.1846 | -34.9947 | -28.8781
18.1907 | -34.9953 | -28.8856
18.1968 | -34.9958 | -28.8932
18.2029 | -34.9964 | -28.9006
18.209 | -34.9969 | -28.9081
18.2151 | -34.9974 | -28.9155
18.2212 | -34.998 | -28.923
18.2273 | -34.9986 | -28.9303
18.2334 | -34.9992 | -28.9379
18.2395 | -34.9997 | -28.945
18.2456 | -35.0003 | -28.9526
18.2517 | -35.0008 | -28.9601
18.2578 | -35.0014 | -28.9669
18.2639 | -35.002 | -28.9744
18.2701 | -35.0025 | -28.9817
18.2762 | -35.0031 | -28.989
18.2823 | -35.0036 | -28.9964
18.2884 | -35.0042 | -29.0038
18.2945 | -35.0048 | -29.011
18.3006 | -35.0053 | -29.0184
18.3067 | -35.0059 | -29.026
18.3128 | -35.0065 | -29.0335
18.3189 | -35.007 | -29.0411
18.325 | -35.0076 | -29.0483
18.3311 | -35.0081 | -29.0557
18.3372 | -35.0087 | -29.0632
18.3433 | -35.0092 | -29.0705
18.3494 | -35.0098 | -29.0779
18.3555 | -35.0104 | -29.0853
18.3616 | -35.011 | -29.0925
18.3677 | -35.0115 | -29.0998
18.3738 | -35.0121 | -29.1074
18.3799 | -35.0126 | -29.1145
18.386 | -35.0132 | -29.1219
18.3921 | -35.0137 | -29.1293
18.3982 | -35.0142 | -29.1363
18.4043 | -35.0147 | -29.1437
18.4105 | -35.0153 | -29.1511
18.4166 | -35.0159 | -29.1586
18.4227 | -35.0164 | -29.1661
18.4288 | -35.0169 | -29.1733
18.4349 | -35.0175 | -29.1805
18.441 | -35.018 | -29.1878
18.4471 | -35.0186 | -29.1951
18.4532 | -35.0191 | -29.2026
18.4593 | -35.0197 | -29.21
18.4654 | -35.0202 | -29.2169
18.4715 | -35.0208 | -29.2245
18.4776 | -35.0213 | -29.2319
18.4837 | -35.0219 | -29.2395
18.4898 | -35.0224 | -29.2469
18.4959 | -35.023 | -29.2543
18.502 | -35.0235 | -29.2617
18.5081 | -35.0241 | -29.2693
18.5142 | -35.0246 | -29.277
18.5203 | -35.0251 | -29.2844
18.5264 | -35.0257 | -29.2915
18.5325 | -35.0263 | -29.2989
18.5386 | -35.0268 | -29.3063
18.5447 | -35.0273 | -29.3137
18.5508 | -35.0279 | -29.3211
18.557 | -35.0284 | -29.3288
18.5631 | -35.0289 | -29.3364
18.5692 | -35.0294 | -29.3442
18.5753 | -35.03 | -29.3515
18.5814 | -35.0306 | -29.3587
18.5875 | -35.0312 | -29.3659
18.5936 | -35.0317 | -29.3732
18.5997 | -35.0322 | -29.3807
18.6058 | -35.0328 | -29.3875
18.6119 | -35.0333 | -29.3951
18.618 | -35.0338 | -29.4025
18.6241 | -35.0344 | -29.4101
18.6302 | -35.0349 | -29.4177
18.6363 | -35.0354 | -29.4251
18.6424 | -35.036 | -29.4322
18.6485 | -35.0365 | -29.4393
18.6546 | -35.037 | -29.4464
18.6607 | -35.0376 | -29.454
18.6668 | -35.0381 | -29.4613
18.6729 | -35.0387 | -29.4686
18.679 | -35.0393 | -29.4756
18.6851 | -35.0398 | -29.4828
18.6912 | -35.0404 | -29.4899
18.6974 | -35.0409 | -29.4972
18.7035 | -35.0414 | -29.5046
18.7096 | -35.0419 | -29.5119
18.7157 | -35.0425 | -29.5193
18.7218 | -35.0431 | -29.5266
18.7279 | -35.0436 | -29.5337
18.734 | -35.0441 | -29.5411
18.7401 | -35.0447 | -29.5483
18.7462 | -35.0452 | -29.5555
18.7523 | -35.0458 | -29.563
18.7584 | -35.0464 | -29.5703
18.7645 | -35.0469 | -29.5775
18.7706 | -35.0475 | -29.5849
18.7767 | -35.048 | -29.5921
18.7828 | -35.0485 | -29.5992
18.7889 | -35.0491 | -29.6062
18.795 | -35.0497 | -29.6138
18.8011 | -35.0502 | -29.621
18.8072 | -35.0508 | -29.6286
18.8133 | -35.0513 | -29.636
18.8194 | -35.0519 | -29.6429
18.8255 | -35.0524 | -29.65
18.8316 | -35.053 | -29.6573
18.8377 | -35.0535 | -29.6645
18.8439 | -35.0541 | -29.6718
18.85 | -35.0546 | -29.6791
18.8561 | -35.0552 | -29.6862
18.8622 | -35.0557 | -29.6935
18.8683 | -35.0563 | -29.7009
18.8744 | -35.0568 | -29.7082
18.8805 | -35.0574 | -29.7155
18.8866 | -35.058 | -29.723
18.8927 | -35.0585 | -29.73
18.8988 | -35.0591 | -29.7375
18.9049 | -35.0596 | -29.7446
18.911 | -35.0602 | -29.7519
18.9171 | -35.0608 | -29.7591
18.9232 | -35.0613 | -29.7658
18.9293 | -35.0619 | -29.7729
18.9354 | -35.0624 | -29.78
18.9415 | -35.063 | -29.7871
18.9476 | -35.0635 | -29.7943
18.9537 | -35.0641 | -29.8017
18.9598 | -35.0646 | -29.8089
18.9659 | -35.0652 | -29.816
18.972 | -35.0657 | -29.8232
18.9781 | -35.0662 | -29.8304
18.9843 | -35.0668 | -29.8378
18.9904 | -35.0673 | -29.8451
18.9965 | -35.0679 | -29.8521
19.0026 | -35.0685 | -29.8594
19.0087 | -35.069 | -29.8671
19.0148 | -35.0696 | -29.8743
19.0209 | -35.0701 | -29.8816
19.027 | -35.0707 | -29.8886
19.0331 | -35.0712 | -29.8955
19.0392 | -35.0718 | -29.9027
19.0453 | -35.0723 | -29.9098
19.0514 | -35.0729 | -29.9171
19.0575 | -35.0734 | -29.9246
19.0636 | -35.0738 | -29.9317
19.0697 | -35.0744 | -29.939
19.0758 | -35.0749 | -29.946
19.0819 | -35.0755 | -29.9531
19.088 | -35.0761 | -29.9605
19.0941 | -35.0765 | -29.9675
19.1002 | -35.0771 | -29.9743
19.1063 | -35.0776 | -29.9816
19.1124 | -35.0782 | -29.9886
19.1185 | -35.0788 | -29.9955
19.1246 | -35.0793 | -30.0025
19.1308 | -35.0799 | -30.0098
19.1369 | -35.0805 | -30.0172
19.143 | -35.081 | -30.0245
19.1491 | -35.0815 | -30.0316
19.1552 | -35.082 | -30.0388
19.1613 | -35.0826 | -30.0459
19.1674 | -35.0831 | -30.0532
19.1735 | -35.0837 | -30.0604
19.1796 | -35.0842 | -30.0677
19.1857 | -35.0848 | -30.0748
19.1918 | -35.0853 | -30.0819
19.1979 | -35.0858 | -30.0894
19.204 | -35.0864 | -30.0965
19.2101 | -35.0869 | -30.1036
19.2162 | -35.0875 | -30.1106
19.2223 | -35.0881 | -30.1177
19.2284 | -35.0887 | -30.1251
19.2345 | -35.0892 | -30.1326
19.2406 | -35.0898 | -30.1396
19.2467 | -35.0903 | -30.1469
19.2528 | -35.0909 | -30.154
19.2589 | -35.0914 | -30.1609
19.265 | -35.092 | -30.1682
19.2712 | -35.0926 | -30.1753
19.2773 | -35.0932 | -30.1821
19.2834 | -35.0937 | -30.1892
19.2895 | -35.0943 | -30.1967
19.2956 | -35.0948 | -30.2043
19.3017 | -35.0953 | -30.2114
19.3078 | -35.0958 | -30.2187
19.3139 | -35.0963 | -30.2255
19.32 | -35.0969 | -30.2326
19.3261 | -35.0974 | -30.2396
19.3322 | -35.098 | -30.2465
19.3383 | -35.0986 | -30.2536
19.3444 | -35.0991 | -30.2603
19.3505 | -35.0996 | -30.2676
19.3566 | -35.1001 | -30.2747
19.3627 | -35.1007 | -30.282
19.3688 | -35.1012 | -30.2889
19.3749 | -35.1018 | -30.2964
19.381 | -35.1023 | -30.3034
19.3871 | -35.1029 | -30.3102
19.3932 | -35.1034 | -30.3175
19.3993 | -35.1039 | -30.3247
19.4054 | -35.1045 | -30.3318
19.4115 | -35.105 | -30.3391
19.4177 | -35.1055 | -30.3465
19.4238 | -35.1059 | -30.3537
19.4299 | -35.1065 | -30.3609
19.436 | -35.1071 | -30.3683
19.4421 | -35.1076 | -30.3757
19.4482 | -35.1081 | -30.383
19.4543 | -35.1087 | -30.3897
19.4604 | -35.1092 | -30.3968
19.4665 | -35.1098 | -30.404
19.4726 | -35.1103 | -30.4112
19.4787 | -35.1109 | -30.4184
19.4848 | -35.1114 | -30.4254
19.4909 | -35.1119 | -30.4327
19.497 | -35.1125 | -30.4397
19.5031 | -35.113 | -30.447
19.5092 | -35.1135 | -30.4541
19.5153 | -35.1141 | -30.4612
19.5214 | -35.1146 | -30.4685
19.5275 | -35.1151 | -30.4752
19.5336 | -35.1156 | -30.4825
19.5397 | -35.1161 | -30.4896
19.5458 | -35.1167 | -30.4966
19.5519 | -35.1172 | -30.5037
19.5581 | -35.1177 | -30.5108
19.5642 | -35.1183 | -30.518
19.5703 | -35.1188 | -30.525
19.5764 | -35.1193 | -30.532
19.5825 | -35.1199 | -30.5391
19.5886 | -35.1204 | -30.5462
19.5947 | -35.121 | -30.5531
19.6008 | -35.1215 | -30.5601
19.6069 | -35.122 | -30.5671
19.613 | -35.1226 | -30.5742
19.6191 | -35.1231 | -30.5812
19.6252 | -35.1236 | -30.5884
19.6313 | -35.1241 | -30.5956
19.6374 | -35.1246 | -30.6027
19.6435 | -35.1251 | -30.6095
19.6496 | -35.1257 | -30.6166
19.6557 | -35.1262 | -30.6234
19.6618 | -35.1267 | -30.6304
19.6679 | -35.1273 | -30.6374
19.674 | -35.1278 | -30.6444
19.6801 | -35.1283 | -30.6518
19.6862 | -35.1289 | -30.659
19.6923 | -35.1294 | -30.6657
19.6984 | -35.1299 | -30.6726
19.7046 | -35.1304 | -30.6796
19.7107 | -35.131 | -30.6864
19.7168 | -35.1315 | -30.6934
19.7229 | -35.1321 | -30.7003
19.729 | -35.1325 | -30.7074
19.7351 | -35.1331 | -30.7141
19.7412 | -35.1336 | -30.7212
19.7473 | -35.1341 | -30.7282
19.7534 | -35.1347 | -30.7353
19.7595 | -35.1351 | -30.7424
19.7656 | -35.1356 | -30.7494
19.7717 | -35.1362 | -30.7557
19.7778 | -35.1367 | -30.7627
19.7839 | -35.1373 | -30.7696
19.79 | -35.1378 | -30.7768
19.7961 | -35.1383 | -30.7841
19.8022 | -35.1387 | -30.7906
19.8083 | -35.1393 | -30.7975
19.8144 | -35.1398 | -30.8045
19.8205 | -35.1404 | -30.8115
19.8266 | -35.1409 | -30.8187
19.8327 | -35.1414 | -30.8257
19.8388 | -35.1419 | -30.8324
19.845 | -35.1425 | -30.8394
19.8511 | -35.143 | -30.8469
19.8572 | -35.1435 | -30.8536
19.8633 | -35.144 | -30.8608
19.8694 | -35.1446 | -30.8683
19.8755 | -35.145 | -30.875
19.8816 | -35.1456 | -30.8819
19.8877 | -35.1461 | -30.8891
19.8938 | -35.1466 | -30.8959
19.8999 | -35.1471 | -30.9032
19.906 | -35.1477 | -30.9101
19.9121 | -35.1482 | -30.917
19.9182 | -35.1487 | -30.924
19.9243 | -35.1492 | -30.931
19.9304 | -35.1497 | -30.9382
19.9365 | -35.1503 | -30.9453
19.9426 | -35.1509 | -30.9526
19.9487 | -35.1514 | -30.9596
19.9548 | -35.1519 | -30.9668
19.9609 | -35.1524 | -30.9741
19.967 | -35.153 | -30.9812
19.9731 | -35.1535 | -30.9884
19.9792 | -35.1541 | -30.9955
19.9853 | -35.1546 | -31.0027
19.9915 | -35.1551 | -31.0097
19.9976 | -35.1556 | -31.017
20.0037 | -35.1561 | -31.0239
20.0098 | -35.1566 | -31.0309
20.0159 | -35.1572 | -31.0378
20.022 | -35.1577 | -31.0451
20.0281 | -35.1582 | -31.0522
20.0342 | -35.1587 | -31.0594
20.0403 | -35.1592 | -31.0666
20.0464 | -35.1597 | -31.0736
20.0525 | -35.1602 | -31.0804
20.0586 | -35.1607 | -31.0873
20.0647 | -35.1612 | -31.0945
20.0708 | -35.1617 | -31.1016
20.0769 | -35.1622 | -31.1087
20.083 | -35.1627 | -31.1154
20.0891 | -35.1632 | -31.1225
20.0952 | -35.1637 | -31.1295
20.1013 | -35.1642 | -31.1368
20.1074 | -35.1648 | -31.144
20.1135 | -35.1653 | -31.1506
20.1196 | -35.1658 | -31.1572
20.1257 | -35.1662 | -31.164
20.1319 | -35.1667 | -31.1711
20.138 | -35.1672 | -31.178
20.1441 | -35.1678 | -31.1854
20.1502 | -35.1682 | -31.1928
20.1563 | -35.1688 | -31.1999
20.1624 | -35.1693 | -31.2066
20.1685 | -35.1699 | -31.2138
20.1746 | -35.1704 | -31.2206
20.1807 | -35.1709 | -31.2277
20.1868 | -35.1714 | -31.2349
20.1929 | -35.1719 | -31.2418
20.199 | -35.1724 | -31.249
20.2051 | -35.1729 | -31.2563
20.2112 | -35.1735 | -31.2633
20.2173 | -35.174 | -31.2701
20.2234 | -35.1745 | -31.2773
20.2295 | -35.1751 | -31.2839
20.2356 | -35.1756 | -31.2909
20.2417 | -35.1761 | -31.2977
20.2478 | -35.1766 | -31.3042
20.2539 | -35.1771 | -31.3111
20.26 | -35.1776 | -31.318
20.2661 | -35.1782 | -31.3253
20.2723 | -35.1788 | -31.3324
20.2784 | -35.1793 | -31.3398
20.2845 | -35.1798 | -31.3468
20.2906 | -35.1803 | -31.3539
20.2967 | -35.1809 | -31.3608
20.3028 | -35.1814 | -31.3677
20.3089 | -35.182 | -31.375
20.315 | -35.1825 | -31.3818
20.3211 | -35.183 | -31.3887
20.3272 | -35.1835 | -31.396
20.3333 | -35.1839 | -31.4026
20.3394 | -35.1844 | -31.4095
20.3455 | -35.1849 | -31.4164
20.3516 | -35.1854 | -31.4231
20.3577 | -35.1859 | -31.4297
20.3638 | -35.1864 | -31.4368
20.3699 | -35.187 | -31.4433
20.376 | -35.1875 | -31.4507
20.3821 | -35.188 | -31.4575
20.3882 | -35.1886 | -31.4645
20.3943 | -35.189 | -31.4715
20.4004 | -35.1895 | -31.4786
20.4065 | -35.19 | -31.4856
20.4126 | -35.1906 | -31.4931
20.4188 | -35.1911 | -31.5
20.4249 | -35.1916 | -31.5067
20.431 | -35.1921 | -31.5135
20.4371 | -35.1926 | -31.5201
20.4432 | -35.1931 | -31.527
20.4493 | -35.1936 | -31.5339
20.4554 | -35.1941 | -31.5405
20.4615 | -35.1946 | -31.5476
20.4676 | -35.1951 | -31.5543
20.4737 | -35.1956 | -31.561
20.4798 | -35.1961 | -31.5681
20.4859 | -35.1967 | -31.5749
20.492 | -35.1973 | -31.5814
20.4981 | -35.1979 | -31.5885
20.5042 | -35.1984 | -31.5958
20.5103 | -35.1989 | -31.6025
20.5164 | -35.1994 | -31.6093
20.5225 | -35.2 | -31.6159
20.5286 | -35.2005 | -31.6228
20.5347 | -35.201 | -31.6295
20.5408 | -35.2016 | -31.6366
20.5469 | -35.2021 | -31.644
20.553 | -35.2027 | -31.6511
20.5592 | -35.2032 | -31.6582
20.5653 | -35.2036 | -31.6651
20.5714 | -35.2042 | -31.6725
20.5775 | -35.2046 | -31.68
20.5836 | -35.2051 | -31.687
20.5897 | -35.2057 | -31.6943
20.5958 | -35.2062 | -31.7018
20.6019 | -35.2068 | -31.7091
20.608 | -35.2073 | -31.7161
20.6141 | -35.2078 | -31.7229
20.6202 | -35.2084 | -31.7301
20.6263 | -35.2088 | -31.7368
20.6324 | -35.2094 | -31.7435
20.6385 | -35.2099 | -31.7502
20.6446 | -35.2104 | -31.7568
20.6507 | -35.211 | -31.7637
20.6568 | -35.2115 | -31.7706
20.6629 | -35.2119 | -31.7773
20.669 | -35.2125 | -31.7842
20.6751 | -35.2129 | -31.7912
20.6812 | -35.2134 | -31.7976
20.6873 | -35.2139 | -31.8043
20.6934 | -35.2144 | -31.8115
20.6995 | -35.2148 | -31.818
20.7057 | -35.2153 | -31.8251
20.7118 | -35.2158 | -31.8324
20.7179 | -35.2163 | -31.839
20.724 | -35.2168 | -31.8463
20.7301 | -35.2173 | -31.8534
20.7362 | -35.2177 | -31.8601
20.7423 | -35.2183 | -31.8663
20.7484 | -35.2188 | -31.8729
20.7545 | -35.2192 | -31.879
20.7606 | -35.2197 | -31.8861
20.7667 | -35.2202 | -31.8933
20.7728 | -35.2207 | -31.8999
20.7789 | -35.2211 | -31.9065
20.785 | -35.2217 | -31.9136
20.7911 | -35.2222 | -31.92
20.7972 | -35.2227 | -31.9273
20.8033 | -35.2233 | -31.9345
20.8094 | -35.2238 | -31.9417
20.8155 | -35.2242 | -31.9485
20.8216 | -35.2248 | -31.9558
20.8277 | -35.2252 | -31.9619
20.8338 | -35.2256 | -31.9685
20.8399 | -35.2261 | -31.975
20.8461 | -35.2267 | -31.9827
20.8522 | -35.2272 | -31.9895
20.8583 | -35.2278 | -31.9961
20.8644 | -35.2283 | -32.0032
20.8705 | -35.2288 | -32.0096
20.8766 | -35.2292 | -32.0166
20.8827 | -35.2298 | -32.0236
20.8888 | -35.2303 | -32.0313
20.8949 | -35.231 | -32.038
20.901 | -35.2315 | -32.0445
20.9071 | -35.2319 | -32.0513
20.9132 | -35.2325 | -32.058
20.9193 | -35.2329 | -32.0647
20.9254 | -35.2334 | -32.0717
20.9315 | -35.234 | -32.0784
20.9376 | -35.2345 | -32.085
20.9437 | -35.235 | -32.0917
20.9498 | -35.2355 | -32.0983
20.9559 | -35.2361 | -32.1053
20.962 | -35.2366 | -32.112
20.9681 | -35.2371 | -32.1186
20.9742 | -35.2376 | -32.1255
20.9803 | -35.2382 | -32.1322
20.9864 | -35.2387 | -32.1388
20.9926 | -35.2392 | -32.1455
20.9987 | -35.2397 | -32.152
21.0048 | -35.2402 | -32.1587
21.0109 | -35.2407 | -32.1649
21.017 | -35.2412 | -32.1715
21.0231 | -35.2418 | -32.1782
21.0292 | -35.2422 | -32.185
21.0353 | -35.2427 | -32.1914
21.0414 | -35.2431 | -32.1979
21.0475 | -35.2436 | -32.205
21.0536 | -35.2441 | -32.2118
21.0597 | -35.2446 | -32.2186
21.0658 | -35.2451 | -32.2255
21.0719 | -35.2456 | -32.2324
21.078 | -35.246 | -32.2396
21.0841 | -35.2465 | -32.2469
21.0902 | -35.247 | -32.2533
21.0963 | -35.2476 | -32.2606
21.1024 | -35.2481 | -32.2679
21.1085 | -35.2487 | -32.2748
21.1146 | -35.2492 | -32.282
21.1207 | -35.2498 | -32.2889
21.1268 | -35.2503 | -32.2962
21.133 | -35.2508 | -32.3026
21.1391 | -35.2515 | -32.309
21.1452 | -35.2519 | -32.3162
21.1513 | -35.2524 | -32.3226
21.1574 | -35.2529 | -32.3294
21.1635 | -35.2534 | -32.336
21.1696 | -35.254 | -32.3425
21.1757 | -35.2546 | -32.3489
21.1818 | -35.2551 | -32.3561
21.1879 | -35.2556 | -32.3628
21.194 | -35.2561 | -32.3697
21.2001 | -35.2566 | -32.3762
21.2062 | -35.2571 | -32.3827
21.2123 | -35.2575 | -32.3895
21.2184 | -35.258 | -32.3963
21.2245 | -35.2585 | -32.4035
21.2306 | -35.259 | -32.41
21.2367 | -35.2596 | -32.4174
21.2428 | -35.2601 | -32.424
21.2489 | -35.2606 | -32.4305
21.255 | -35.2611 | -32.4377
21.2611 | -35.2615 | -32.4445
21.2672 | -35.262 | -32.4512
21.2733 | -35.2625 | -32.4579
21.2795 | -35.2631 | -32.4645
21.2856 | -35.2636 | -32.471
21.2917 | -35.2641 | -32.4774
21.2978 | -35.2645 | -32.4839
21.3039 | -35.265 | -32.491
21.31 | -35.2655 | -32.4978
21.3161 | -35.2659 | -32.504
21.3222 | -35.2663 | -32.5104
21.3283 | -35.2667 | -32.5168
21.3344 | -35.2673 | -32.5234
21.3405 | -35.2678 | -32.5301
21.3466 | -35.2683 | -32.5367
21.3527 | -35.2688 | -32.5439
21.3588 | -35.2693 | -32.551
21.3649 | -35.2699 | -32.5574
21.371 | -35.2704 | -32.5642
21.3771 | -35.2709 | -32.571
21.3832 | -35.2713 | -32.5775
21.3893 | -35.2718 | -32.5838
21.3954 | -35.2723 | -32.5911
21.4015 | -35.2728 | -32.5975
21.4076 | -35.2732 | -32.604
21.4137 | -35.2737 | -32.6112
21.4199 | -35.2742 | -32.6176
21.426 | -35.2748 | -32.6236
21.4321 | -35.2752 | -32.6299
21.4382 | -35.2757 | -32.6369
21.4443 | -35.2761 | -32.6432
21.4504 | -35.2765 | -32.6498
21.4565 | -35.2771 | -32.656
21.4626 | -35.2775 | -32.6627
21.4687 | -35.2781 | -32.6693
21.4748 | -35.2786 | -32.6758
21.4809 | -35.2791 | -32.6827
21.487 | -35.2795 | -32.6895
21.4931 | -35.2799 | -32.6959
21.4992 | -35.2805 | -32.7029
21.5053 | -35.281 | -32.7092
21.5114 | -35.2815 | -32.716
21.5175 | -35.282 | -32.723
21.5236 | -35.2825 | -32.7292
21.5297 | -35.2831 | -32.7355
21.5358 | -35.2835 | -32.7422
21.5419 | -35.284 | -32.7492
21.548 | -35.2845 | -32.7555
21.5541 | -35.2849 | -32.7625
21.5602 | -35.2854 | -32.7691
21.5664 | -35.2859 | -32.776
21.5725 | -35.2865 | -32.7829
21.5786 | -35.2871 | -32.7896
21.5847 | -35.2876 | -32.7972
21.5908 | -35.2881 | -32.8037
21.5969 | -35.2887 | -32.8102
21.603 | -35.2891 | -32.8172
21.6091 | -35.2896 | -32.8245
21.6152 | -35.2901 | -32.8312
21.6213 | -35.2906 | -32.8381
21.6274 | -35.2912 | -32.8447
21.6335 | -35.2915 | -32.8512
21.6396 | -35.2921 | -32.8581
21.6457 | -35.2926 | -32.8646
21.6518 | -35.293 | -32.8707
21.6579 | -35.2935 | -32.8772
21.664 | -35.294 | -32.8834
21.6701 | -35.2945 | -32.8901
21.6762 | -35.2949 | -32.8966
21.6823 | -35.2954 | -32.9033
21.6884 | -35.2959 | -32.91
21.6945 | -35.2964 | -32.917
21.7006 | -35.2968 | -32.9234
21.7068 | -35.2973 | -32.9299
21.7129 | -35.2978 | -32.9365
21.719 | -35.2983 | -32.9433
21.7251 | -35.2988 | -32.9503
21.7312 | -35.2993 | -32.9569
21.7373 | -35.2999 | -32.9626
21.7434 | -35.3004 | -32.9687
21.7495 | -35.3009 | -32.9756
21.7556 | -35.3015 | -32.9828
21.7617 | -35.3019 | -32.9894
21.7678 | -35.3024 | -32.9955
21.7739 | -35.3029 | -33.0012
21.78 | -35.3033 | -33.0078
21.7861 | -35.3038 | -33.0144
21.7922 | -35.3043 | -33.0216
21.7983 | -35.3049 | -33.0289
21.8044 | -35.3053 | -33.0357
21.8105 | -35.3058 | -33.0414
21.8166 | -35.3063 | -33.047
21.8227 | -35.3069 | -33.0531
21.8288 | -35.3074 | -33.0594
21.8349 | -35.3078 | -33.0666
21.841 | -35.3083 | -33.0725
21.8471 | -35.3087 | -33.0781
21.8533 | -35.3092 | -33.0851
21.8594 | -35.3097 | -33.0915
21.8655 | -35.3102 | -33.0982
21.8716 | -35.3107 | -33.105
21.8777 | -35.3112 | -33.1118
21.8838 | -35.3117 | -33.1184
21.8899 | -35.3122 | -33.1252
21.896 | -35.3127 | -33.1317
21.9021 | -35.3132 | -33.1385
21.9082 | -35.3137 | -33.1448
21.9143 | -35.3142 | -33.15
21.9204 | -35.3147 | -33.1558
21.9265 | -35.3152 | -33.1629
21.9326 | -35.3157 | -33.1687
21.9387 | -35.3162 | -33.176
21.9448 | -35.3167 | -33.1828
21.9509 | -35.3173 | -33.1884
21.957 | -35.3177 | -33.1957
21.9631 | -35.3182 | -33.2018
21.9692 | -35.3186 | -33.2089
21.9753 | -35.3191 | -33.2157
21.9814 | -35.3197 | -33.222
21.9875 | -35.3203 | -33.2274
21.9937 | -35.3208 | -33.2337
21.9998 | -35.3214 | -33.2413
22.0059 | -35.3219 | -33.2478
22.012 | -35.3222 | -33.2544
22.0181 | -35.3227 | -33.2616
22.0242 | -35.3232 | -33.2683
22.0303 | -35.3239 | -33.2754
22.0364 | -35.3243 | -33.2819
22.0425 | -35.325 | -33.2883
22.0486 | -35.3256 | -33.2949
22.0547 | -35.3262 | -33.3
22.0608 | -35.3266 | -33.3061
22.0669 | -35.3271 | -33.3134
22.073 | -35.3275 | -33.3203
22.0791 | -35.3281 | -33.3263
22.0852 | -35.3285 | -33.3336
22.0913 | -35.3289 | -33.3385
22.0974 | -35.3295 | -33.3455
22.1035 | -35.3298 | -33.3521
22.1096 | -35.3303 | -33.3583
22.1157 | -35.3308 | -33.3645
22.1218 | -35.3313 | -33.3708
22.1279 | -35.3318 | -33.377
22.134 | -35.3323 | -33.3837
22.1402 | -35.3328 | -33.3897
22.1463 | -35.3332 | -33.3962
22.1524 | -35.3338 | -33.403
22.1585 | -35.3343 | -33.4097
22.1646 | -35.3348 | -33.4159
22.1707 | -35.3353 | -33.4232
22.1768 | -35.3358 | -33.4298
22.1829 | -35.3363 | -33.4363
22.189 | -35.3368 | -33.4432
22.1951 | -35.3373 | -33.449
22.2012 | -35.338 | -33.456
22.2073 | -35.3385 | -33.4627
22.2134 | -35.339 | -33.4697
22.2195 | -35.3396 | -33.4769
22.2256 | -35.3402 | -33.4832
22.2317 | -35.3406 | -33.49
22.2378 | -35.3411 | -33.4959
22.2439 | -35.3417 | -33.5015
22.25 | -35.3422 | -33.5079
22.2561 | -35.3428 | -33.5146
22.2622 | -35.3433 | -33.5208
22.2683 | -35.3438 | -33.527
22.2744 | -35.3442 | -33.5339
22.2806 | -35.3447 | -33.5404
22.2867 | -35.3452 | -33.5468
22.2928 | -35.3457 | -33.5531
22.2989 | -35.3463 | -33.56
22.305 | -35.3468 | -33.5663
22.3111 | -35.3473 | -33.5732
22.3172 | -35.3478 | -33.58
22.3233 | -35.3483 | -33.5862
22.3294 | -35.3488 | -33.5927
22.3355 | -35.3494 | -33.5992
22.3416 | -35.35 | -33.606
22.3477 | -35.3504 | -33.6123
22.3538 | -35.3509 | -33.6186
22.3599 | -35.3514 | -33.6248
22.366 | -35.3519 | -33.6312
22.3721 | -35.3524 | -33.6378
22.3782 | -35.3529 | -33.6446
22.3843 | -35.3535 | -33.6509
22.3904 | -35.3539 | -33.6581
22.3965 | -35.3544 | -33.6648
22.4026 | -35.3548 | -33.6716
22.4087 | -35.3553 | -33.6789
22.4148 | -35.3558 | -33.6847
22.4209 | -35.3563 | -33.6913
22.4271 | -35.3568 | -33.6981
22.4332 | -35.3573 | -33.7051
22.4393 | -35.3579 | -33.7112
22.4454 | -35.3584 | -33.7175
22.4515 | -35.359 | -33.724
22.4576 | -35.3596 | -33.73
22.4637 | -35.36 | -33.7363
22.4698 | -35.3604 | -33.7423
22.4759 | -35.3609 | -33.7493
22.482 | -35.3614 | -33.756
22.4881 | -35.3619 | -33.7623
22.4942 | -35.3624 | -33.7687
22.5003 | -35.3629 | -33.7756
22.5064 | -35.3634 | -33.7819
22.5125 | -35.3639 | -33.7883
22.5186 | -35.3643 | -33.7943
22.5247 | -35.3648 | -33.801
22.5308 | -35.3654 | -33.8073
22.5369 | -35.3659 | -33.8139
22.543 | -35.3664 | -33.8204
22.5491 | -35.3669 | -33.8273
22.5552 | -35.3674 | -33.8342
22.5613 | -35.368 | -33.8406
22.5675 | -35.3685 | -33.8467
22.5736 | -35.3689 | -33.8533
22.5797 | -35.3695 | -33.8592
22.5858 | -35.3699 | -33.8649
22.5919 | -35.3705 | -33.8717
22.598 | -35.371 | -33.8777
22.6041 | -35.3716 | -33.8843
22.6102 | -35.3721 | -33.8909
22.6163 | -35.3725 | -33.897
22.6224 | -35.373 | -33.9034
22.6285 | -35.3735 | -33.9103
22.6346 | -35.374 | -33.917
22.6407 | -35.3746 | -33.9234
22.6468 | -35.3752 | -33.9296
22.6529 | -35.3756 | -33.9356
22.659 | -35.3761 | -33.942
22.6651 | -35.3765 | -33.9485
22.6712 | -35.377 | -33.9545
22.6773 | -35.3775 | -33.9608
22.6834 | -35.378 | -33.9669
22.6895 | -35.3785 | -33.9727
22.6956 | -35.3791 | -33.9792
22.7017 | -35.3796 | -33.9852
22.7079 | -35.38 | -33.9908
22.714 | -35.3805 | -33.9971
22.7201 | -35.3808 | -34.0036
22.7262 | -35.3813 | -34.0096
22.7323 | -35.3818 | -34.0162
22.7384 | -35.3821 | -34.0229
22.7445 | -35.3826 | -34.0292
22.7506 | -35.3832 | -34.0355
22.7567 | -35.3837 | -34.0417
22.7628 | -35.3845 | -34.0488
22.7689 | -35.3851 | -34.0549
22.775 | -35.3856 | -34.0605
22.7811 | -35.386 | -34.0673
22.7872 | -35.3865 | -34.0741
22.7933 | -35.387 | -34.0801
22.7994 | -35.3876 | -34.0863
22.8055 | -35.3881 | -34.0923
22.8116 | -35.3886 | -34.0989
22.8177 | -35.3891 | -34.1059
22.8238 | -35.3897 | -34.1121
22.8299 | -35.3903 | -34.1184
22.836 | -35.3908 | -34.1249
22.8421 | -35.3914 | -34.131
22.8482 | -35.3918 | -34.1372
22.8544 | -35.3923 | -34.1441
22.8605 | -35.3928 | -34.1503
22.8666 | -35.3934 | -34.1573
22.8727 | -35.394 | -34.1635
22.8788 | -35.3945 | -34.1697
22.8849 | -35.395 | -34.1762
22.891 | -35.3955 | -34.1823
22.8971 | -35.3961 | -34.1886
22.9032 | -35.3967 | -34.1944
22.9093 | -35.3972 | -34.2005
22.9154 | -35.3978 | -34.2066
22.9215 | -35.3982 | -34.2132
22.9276 | -35.3987 | -34.2198
22.9337 | -35.3992 | -34.2258
22.9398 | -35.3998 | -34.2313
22.9459 | -35.4003 | -34.2385
22.952 | -35.4008 | -34.2443
22.9581 | -35.4012 | -34.2502
22.9642 | -35.4017 | -34.2556
22.9703 | -35.4023 | -34.2603
22.9764 | -35.4028 | -34.2668
22.9825 | -35.4032 | -34.2734
22.9886 | -35.4037 | -34.2797
22.9948 | -35.4041 | -34.286
23.0009 | -35.4046 | -34.2923
23.007 | -35.4051 | -34.2983
23.0131 | -35.4056 | -34.3039
23.0192 | -35.4062 | -34.3101
23.0253 | -35.4067 | -34.3162
23.0314 | -35.4072 | -34.3224
23.0375 | -35.4077 | -34.329
23.0436 | -35.4082 | -34.3354
23.0497 | -35.4088 | -34.3417
23.0558 | -35.4093 | -34.3484
23.0619 | -35.4098 | -34.3547
23.068 | -35.4103 | -34.3606
23.0741 | -35.4109 | -34.3669
23.0802 | -35.4113 | -34.3735
23.0863 | -35.4119 | -34.3795
23.0924 | -35.4124 | -34.3856
23.0985 | -35.4129 | -34.3918
23.1046 | -35.4134 | -34.3985
23.1107 | -35.4139 | -34.4041
23.1168 | -35.4145 | -34.4103
23.1229 | -35.415 | -34.4161
23.129 | -35.4155 | -34.4226
23.1351 | -35.4161 | -34.4291
23.1413 | -35.4165 | -34.4352
23.1474 | -35.417 | -34.4409
23.1535 | -35.4175 | -34.447
23.1596 | -35.4179 | -34.4537
23.1657 | -35.4185 | -34.46
23.1718 | -35.419 | -34.466
23.1779 | -35.4196 | -34.4722
23.184 | -35.4201 | -34.4792
23.1901 | -35.4206 | -34.486
23.1962 | -35.4212 | -34.4918
23.2023 | -35.4217 | -34.4985
23.2084 | -35.4222 | -34.505
23.2145 | -35.4226 | -34.5114
23.2206 | -35.4233 | -34.5176
23.2267 | -35.4238 | -34.5238
23.2328 | -35.4242 | -34.5296
23.2389 | -35.4247 | -34.5358
23.245 | -35.4253 | -34.5427
23.2511 | -35.4259 | -34.5482
23.2572 | -35.4265 | -34.554
23.2633 | -35.427 | -34.5602
23.2694 | -35.4275 | -34.566
23.2755 | -35.428 | -34.5727
23.2817 | -35.4286 | -34.5782
23.2878 | -35.4291 | -34.5835
23.2939 | -35.4296 | -34.5898
23.3 | -35.4301 | -34.5965
23.3061 | -35.4306 | -34.6025
23.3122 | -35.4311 | -34.6085
23.3183 | -35.4316 | -34.6147
23.3244 | -35.4321 | -34.6207
23.3305 | -35.4325 | -34.6268
23.3366 | -35.433 | -34.6331
23.3427 | -35.4336 | -34.6395
23.3488 | -35.4341 | -34.6459
23.3549 | -35.4345 | -34.6518
23.361 | -35.435 | -34.6572
23.3671 | -35.4355 | -34.6636
23.3732 | -35.436 | -34.6702
23.3793 | -35.4365 | -34.6765
23.3854 | -35.4371 | -34.683
23.3915 | -35.4376 | -34.6888
23.3976 | -35.4381 | -34.695
23.4037 | -35.4387 | -34.7006
23.4098 | -35.4392 | -34.7068
23.4159 | -35.4397 | -34.7126
23.422 | -35.4402 | -34.7182
23.4282 | -35.4408 | -34.7236
23.4343 | -35.4413 | -34.7301
23.4404 | -35.4419 | -34.7369
23.4465 | -35.4424 | -34.7427
23.4526 | -35.4429 | -34.7495
23.4587 | -35.4435 | -34.7557
23.4648 | -35.4438 | -34.7626
23.4709 | -35.4442 | -34.7682
23.477 | -35.4448 | -34.7737
23.4831 | -35.4453 | -34.7804
23.4892 | -35.4459 | -34.7864
23.4953 | -35.4464 | -34.7919
23.5014 | -35.4471 | -34.797
23.5075 | -35.4475 | -34.8036
23.5136 | -35.4481 | -34.8096
23.5197 | -35.4485 | -34.8157
23.5258 | -35.4489 | -34.822
23.5319 | -35.4495 | -34.8277
23.538 | -35.4501 | -34.8338
23.5441 | -35.4506 | -34.8402
23.5502 | -35.4511 | -34.8464
23.5563 | -35.4516 | -34.8527
23.5624 | -35.4521 | -34.8582
23.5686 | -35.4526 | -34.864
23.5747 | -35.4531 | -34.8694
23.5808 | -35.4535 | -34.8757
23.5869 | -35.454 | -34.8815
23.593 | -35.4545 | -34.8877
23.5991 | -35.455 | -34.8942
23.6052 | -35.4554 | -34.8999
23.6113 | -35.456 | -34.9065
23.6174 | -35.4564 | -34.913
23.6235 | -35.457 | -34.9186
23.6296 | -35.4575 | -34.9254
23.6357 | -35.458 | -34.9307
23.6418 | -35.4585 | -34.9367
23.6479 | -35.459 | -34.9433
23.654 | -35.4595 | -34.9488
23.6601 | -35.4601 | -34.9542
23.6662 | -35.4606 | -34.9603
23.6723 | -35.4611 | -34.9656
23.6784 | -35.4617 | -34.9713
23.6845 | -35.4623 | -34.9776
23.6906 | -35.4626 | -34.9839
23.6967 | -35.4632 | -34.99
23.7028 | -35.4638 | -34.9961
23.7089 | -35.4644 | -35.0025
23.7151 | -35.4649 | -35.008
23.7212 | -35.4654 | -35.014
23.7273 | -35.4659 | -35.0203
23.7334 | -35.4664 | -35.0262
23.7395 | -35.4669 | -35.0327
23.7456 | -35.4674 | -35.0382
23.7517 | -35.4679 | -35.0442
23.7578 | -35.4684 | -35.0496
23.7639 | -35.469 | -35.0556
23.77 | -35.4695 | -35.0619
23.7761 | -35.4701 | -35.0677
23.7822 | -35.4706 | -35.0729
23.7883 | -35.4711 | -35.0787
23.7944 | -35.4718 | -35.0847
23.8005 | -35.4723 | -35.091
23.8066 | -35.4728 | -35.097
23.8127 | -35.4732 | -35.1031
23.8188 | -35.4738 | -35.1088
23.8249 | -35.4743 | -35.1145
23.831 | -35.4748 | -35.121
23.8371 | -35.4754 | -35.1273
23.8432 | -35.4759 | -35.1331
23.8493 | -35.4765 | -35.1396
23.8555 | -35.477 | -35.1465
23.8616 | -35.4775 | -35.1515
23.8677 | -35.478 | -35.1575
23.8738 | -35.4785 | -35.1637
23.8799 | -35.4791 | -35.1692
23.886 | -35.4796 | -35.1738
23.8921 | -35.4802 | -35.1795
23.8982 | -35.4807 | -35.1856
23.9043 | -35.4812 | -35.1922
23.9104 | -35.4817 | -35.1979
23.9165 | -35.4822 | -35.2041
23.9226 | -35.4827 | -35.2106
23.9287 | -35.4833 | -35.2169
23.9348 | -35.4838 | -35.2225
23.9409 | -35.4843 | -35.2285
23.947 | -35.4848 | -35.2342
23.9531 | -35.4853 | -35.2403
23.9592 | -35.4859 | -35.2465
23.9653 | -35.4864 | -35.2526
23.9714 | -35.4869 | -35.2585
23.9775 | -35.4874 | -35.2643
23.9836 | -35.4879 | -35.2708
23.9897 | -35.4885 | -35.2758
23.9958 | -35.489 | -35.2816
24.002 | -35.4895 | -35.2867
24.0081 | -35.4901 | -35.2929
24.0142 | -35.4905 | -35.2982
24.0203 | -35.491 | -35.3044
24.0264 | -35.4915 | -35.3101
24.0325 | -35.492 | -35.3158
24.0386 | -35.4925 | -35.3219
24.0447 | -35.493 | -35.3278
24.0508 | -35.4936 | -35.3341
24.0569 | -35.4941 | -35.34
24.063 | -35.4947 | -35.346
24.0691 | -35.4952 | -35.3528
24.0752 | -35.4958 | -35.3591
24.0813 | -35.4964 | -35.3657
24.0874 | -35.4969 | -35.3717
24.0935 | -35.4974 | -35.3773
24.0996 | -35.4979 | -35.3824
24.1057 | -35.4985 | -35.3885
24.1118 | -35.499 | -35.3944
24.1179 | -35.4996 | -35.4008
24.124 | -35.5002 | -35.4067
24.1301 | -35.5007 | -35.412
24.1362 | -35.5013 | -35.4181
24.1424 | -35.5018 | -35.4246
24.1485 | -35.5023 | -35.4305
24.1546 | -35.5028 | -35.4359
24.1607 | -35.5033 | -35.4413
24.1668 | -35.5039 | -35.4464
24.1729 | -35.5045 | -35.4525
24.179 | -35.505 | -35.4582
24.1851 | -35.5054 | -35.4642
24.1912 | -35.5059 | -35.4697
24.1973 | -35.5064 | -35.4752
24.2034 | -35.507 | -35.4804
24.2095 | -35.5076 | -35.4864
24.2156 | -35.508 | -35.4922
24.2217 | -35.5085 | -35.4988
24.2278 | -35.509 | -35.5051
24.2339 | -35.5095 | -35.5106
24.24 | -35.5101 | -35.5155
24.2461 | -35.5106 | -35.5212
24.2522 | -35.5111 | -35.5269
24.2583 | -35.5117 | -35.533
24.2644 | -35.5122 | -35.5388
24.2705 | -35.5127 | -35.5447
24.2766 | -35.5133 | -35.5507
24.2827 | -35.5138 | -35.5571
24.2889 | -35.5143 | -35.5628
24.295 | -35.5148 | -35.5689
24.3011 | -35.5153 | -35.5747
24.3072 | -35.5159 | -35.5809
24.3133 | -35.5164 | -35.5868
24.3194 | -35.5169 | -35.5928
24.3255 | -35.5174 | -35.599
24.3316 | -35.5179 | -35.6046
24.3377 | -35.5185 | -35.6106
24.3438 | -35.5191 | -35.616
24.3499 | -35.5196 | -35.6218
24.356 | -35.5201 | -35.6274
24.3621 | -35.5205 | -35.6329
24.3682 | -35.5211 | -35.6383
24.3743 | -35.5216 | -35.6443
24.3804 | -35.5222 | -35.6496
24.3865 | -35.5227 | -35.6555
24.3926 | -35.5232 | -35.6614
24.3987 | -35.5237 | -35.6669
24.4048 | -35.5242 | -35.6728
24.4109 | -35.5248 | -35.6787
24.417 | -35.5254 | -35.6847
24.4231 | -35.5259 | -35.6905
24.4293 | -35.5263 | -35.6962
24.4354 | -35.5268 | -35.7016
24.4415 | -35.5275 | -35.7075
24.4476 | -35.5279 | -35.7131
24.4537 | -35.5284 | -35.7187
24.4598 | -35.529 | -35.7249
24.4659 | -35.5295 | -35.7302
24.472 | -35.5301 | -35.7363
24.4781 | -35.5306 | -35.7418
24.4842 | -35.5312 | -35.7466
24.4903 | -35.5317 | -35.7529
24.4964 | -35.5322 | -35.7582
24.5025 | -35.5328 | -35.7641
24.5086 | -35.5333 | -35.7697
24.5147 | -35.5339 | -35.7761
24.5208 | -35.5343 | -35.7816
24.5269 | -35.5348 | -35.787
24.533 | -35.5355 | -35.7931
24.5391 | -35.5361 | -35.7991
24.5452 | -35.5367 | -35.8048
24.5513 | -35.5371 | -35.8108
24.5574 | -35.5377 | -35.8174
24.5635 | -35.5381 | -35.8228
24.5696 | -35.5387 | -35.8285
24.5758 | -35.5391 | -35.8346
24.5819 | -35.5397 | -35.8397
24.588 | -35.5402 | -35.8441
24.5941 | -35.5406 | -35.8499
24.6002 | -35.5411 | -35.8553
24.6063 | -35.5417 | -35.8613
24.6124 | -35.5422 | -35.8669
24.6185 | -35.5428 | -35.8729
24.6246 | -35.5435 | -35.8785
24.6307 | -35.5439 | -35.8849
24.6368 | -35.5445 | -35.8904
24.6429 | -35.545 | -35.8959
24.649 | -35.5455 | -35.9015
24.6551 | -35.546 | -35.9071
24.6612 | -35.5464 | -35.9124
24.6673 | -35.547 | -35.9185
24.6734 | -35.5474 | -35.9241
24.6795 | -35.548 | -35.9299
24.6856 | -35.5486 | -35.9359
24.6917 | -35.5492 | -35.9417
24.6978 | -35.5497 | -35.9476
24.7039 | -35.5501 | -35.9535
24.71 | -35.5506 | -35.9594
24.7162 | -35.5511 | -35.9653
24.7223 | -35.5516 | -35.9705
24.7284 | -35.5521 | -35.9753
24.7345 | -35.5527 | -35.9817
24.7406 | -35.5532 | -35.9872
24.7467 | -35.5537 | -35.9934
24.7528 | -35.5543 | -35.9997
24.7589 | -35.5547 | -36.0053
24.765 | -35.5552 | -36.01
24.7711 | -35.5557 | -36.0159
24.7772 | -35.5562 | -36.0214
24.7833 | -35.5568 | -36.0273
24.7894 | -35.5574 | -36.033
24.7955 | -35.5578 | -36.039
24.8016 | -35.5582 | -36.0449
24.8077 | -35.5587 | -36.0514
24.8138 | -35.5593 | -36.057
24.8199 | -35.5598 | -36.0622
24.826 | -35.5603 | -36.0678
24.8321 | -35.5609 | -36.074
24.8382 | -35.5615 | -36.0794
24.8443 | -35.5619 | -36.085
24.8504 | -35.5624 | -36.0912
24.8565 | -35.5629 | -36.098
24.8627 | -35.5635 | -36.1037
24.8688 | -35.5639 | -36.1102
24.8749 | -35.5644 | -36.1164
24.881 | -35.5649 | -36.122
24.8871 | -35.5654 | -36.1278
24.8932 | -35.566 | -36.1337
24.8993 | -35.5667 | -36.1393
24.9054 | -35.5672 | -36.145
24.9115 | -35.5678 | -36.151
24.9176 | -35.5682 | -36.1575
24.9237 | -35.5687 | -36.1624
24.9298 | -35.5692 | -36.1684
24.9359 | -35.5697 | -36.1738
24.942 | -35.5703 | -36.1788
24.9481 | -35.5708 | -36.1845
24.9542 | -35.5713 | -36.1902
24.9603 | -35.5718 | -36.1954
24.9664 | -35.5723 | -36.2015
24.9725 | -35.5729 | -36.2079
24.9786 | -35.5733 | -36.2132
24.9847 | -35.5738 | -36.2187
24.9908 | -35.5743 | -36.2245
24.9969 | -35.5747 | -36.2303
25.0031 | -35.5753 | -36.2356
25.0092 | -35.5759 | -36.2405
25.0153 | -35.5764 | -36.2463
25.0214 | -35.5769 | -36.253
25.0275 | -35.5774 | -36.2589
25.0336 | -35.5779 | -36.2643
25.0397 | -35.5784 | -36.2699
25.0458 | -35.5789 | -36.2764
25.0519 | -35.5794 | -36.2815
25.058 | -35.5798 | -36.287
25.0641 | -35.5804 | -36.2922
25.0702 | -35.581 | -36.2985
25.0763 | -35.5815 | -36.3046
25.0824 | -35.5818 | -36.3101
25.0885 | -35.5822 | -36.3157
25.0946 | -35.5827 | -36.3216
25.1007 | -35.5833 | -36.3282
25.1068 | -35.5839 | -36.3334
25.1129 | -35.5844 | -36.3394
25.119 | -35.585 | -36.3447
25.1251 | -35.5856 | -36.3507
25.1312 | -35.5861 | -36.3562
25.1373 | -35.5867 | -36.3626
25.1435 | -35.5872 | -36.3687
25.1496 | -35.5877 | -36.3752
25.1557 | -35.5882 | -36.3803
25.1618 | -35.5887 | -36.3857
25.1679 | -35.5893 | -36.391
25.174 | -35.5899 | -36.3967
25.1801 | -35.5904 | -36.4029
25.1862 | -35.5908 | -36.4093
25.1923 | -35.5914 | -36.4141
25.1984 | -35.592 | -36.4197
25.2045 | -35.5924 | -36.4252
25.2106 | -35.5929 | -36.4307
25.2167 | -35.5934 | -36.436
25.2228 | -35.5939 | -36.4425
25.2289 | -35.5943 | -36.4484
25.235 | -35.5949 | -36.454
25.2411 | -35.5953 | -36.4594
25.2472 | -35.5957 | -36.4649
25.2533 | -35.5962 | -36.4705
25.2594 | -35.5967 | -36.4759
25.2655 | -35.5973 | -36.482
25.2716 | -35.5978 | -36.487
25.2777 | -35.5983 | -36.4933
25.2838 | -35.5988 | -36.4986
25.29 | -35.5992 | -36.5044
25.2961 | -35.5997 | -36.5105
25.3022 | -35.6001 | -36.516
25.3083 | -35.6006 | -36.5218
25.3144 | -35.6012 | -36.5276
25.3205 | -35.6018 | -36.5341
25.3266 | -35.6024 | -36.5403
25.3327 | -35.6029 | -36.5454
25.3388 | -35.6034 | -36.5509
25.3449 | -35.6039 | -36.5572
25.351 | -35.6045 | -36.5622
25.3571 | -35.6049 | -36.5685
25.3632 | -35.6056 | -36.5736
25.3693 | -35.6061 | -36.579
25.3754 | -35.6066 | -36.5841
25.3815 | -35.607 | -36.5902
25.3876 | -35.6075 | -36.5961
25.3937 | -35.608 | -36.6021
25.3998 | -35.6085 | -36.6074
25.4059 | -35.609 | -36.6131
25.412 | -35.6095 | -36.6186
25.4181 | -35.6101 | -36.6244
25.4242 | -35.6105 | -36.6298
25.4304 | -35.6111 | -36.6353
25.4365 | -35.6115 | -36.6421
25.4426 | -35.6121 | -36.6475
25.4487 | -35.6126 | -36.653
25.4548 | -35.6131 | -36.6587
25.4609 | -35.6137 | -36.6639
25.467 | -35.6141 | -36.6689
25.4731 | -35.6146 | -36.6749
25.4792 | -35.6151 | -36.6805
25.4853 | -35.6157 | -36.686
25.4914 | -35.6162 | -36.6916
25.4975 | -35.6167 | -36.6972
25.5036 | -35.6171 | -36.7025
25.5097 | -35.6175 | -36.7078
25.5158 | -35.618 | -36.7135
25.5219 | -35.6184 | -36.7187
25.528 | -35.6188 | -36.7237
25.5341 | -35.6193 | -36.7291
25.5402 | -35.6197 | -36.735
25.5463 | -35.6203 | -36.7402
25.5524 | -35.6208 | -36.746
25.5585 | -35.6213 | -36.7517
25.5646 | -35.6218 | -36.7574
25.5707 | -35.6223 | -36.7623
25.5769 | -35.6227 | -36.7677
25.583 | -35.6231 | -36.7731
25.5891 | -35.6236 | -36.7786
25.5952 | -35.6241 | -36.7845
25.6013 | -35.6246 | -36.7904
25.6074 | -35.6251 | -36.7945
25.6135 | -35.6256 | -36.8001
25.6196 | -35.6262 | -36.8058
25.6257 | -35.6266 | -36.8114
25.6318 | -35.6271 | -36.8173
25.6379 | -35.6276 | -36.823
25.644 | -35.6281 | -36.8281
25.6501 | -35.6286 | -36.8342
25.6562 | -35.629 | -36.8401
25.6623 | -35.6294 | -36.845
25.6684 | -35.6298 | -36.8505
25.6745 | -35.6304 | -36.8564
25.6806 | -35.631 | -36.8617
25.6867 | -35.6315 | -36.867
25.6928 | -35.632 | -36.8728
25.6989 | -35.6324 | -36.8781
25.705 | -35.6329 | -36.8839
25.7111 | -35.6334 | -36.8897
25.7173 | -35.6338 | -36.8948
25.7234 | -35.6343 | -36.9005
25.7295 | -35.6348 | -36.9058
25.7356 | -35.6352 | -36.9115
25.7417 | -35.6357 | -36.9172
25.7478 | -35.6362 | -36.9223
25.7539 | -35.6367 | -36.9279
25.76 | -35.6372 | -36.9333
25.7661 | -35.6376 | -36.9391
25.7722 | -35.6382 | -36.9449
25.7783 | -35.6387 | -36.9512
25.7844 | -35.6391 | -36.9573
25.7905 | -35.6395 | -36.9628
25.7966 | -35.64 | -36.9689
25.8027 | -35.6405 | -36.9742
25.8088 | -35.641 | -36.9797
25.8149 | -35.6415 | -36.9848
25.821 | -35.642 | -36.991
25.8271 | -35.6425 | -36.9965
25.8332 | -35.6429 | -37.0023
25.8393 | -35.6434 | -37.0083
25.8454 | -35.6438 | -37.0137
25.8515 | -35.6443 | -37.0196
25.8576 | -35.6448 | -37.025
25.8638 | -35.6452 | -37.0304
25.8699 | -35.6457 | -37.0364
25.876 | -35.6463 | -37.0418
25.8821 | -35.6467 | -37.0471
25.8882 | -35.6471 | -37.0534
25.8943 | -35.6476 | -37.0594
25.9004 | -35.648 | -37.0652
25.9065 | -35.6484 | -37.071
25.9126 | -35.6489 | -37.0766
25.9187 | -35.6494 | -37.0818
25.9248 | -35.6499 | -37.0878
25.9309 | -35.6503 | -37.0937
25.937 | -35.6508 | -37.0997
25.9431 | -35.6512 | -37.1057
25.9492 | -35.6516 | -37.1113
25.9553 | -35.6521 | -37.1173
25.9614 | -35.6526 | -37.1227
25.9675 | -35.6531 | -37.1287
25.9736 | -35.6536 | -37.1343
25.9797 | -35.6541 | -37.1406
25.9858 | -35.6545 | -37.1462
25.9919 | -35.655 | -37.1522
25.998 | -35.6554 | -37.1576
26.0042 | -35.656 | -37.1633
26.0103 | -35.6564 | -37.1687
26.0164 | -35.6569 | -37.1745
26.0225 | -35.6573 | -37.1802
26.0286 | -35.6578 | -37.1858
26.0347 | -35.6583 | -37.1914
26.0408 | -35.6587 | -37.1969
26.0469 | -35.6591 | -37.2024
26.053 | -35.6596 | -37.2078
26.0591 | -35.6601 | -37.2137
26.0652 | -35.6605 | -37.2188
26.0713 | -35.661 | -37.2243
26.0774 | -35.6614 | -37.23
26.0835 | -35.6619 | -37.2352
26.0896 | -35.6624 | -37.2411
26.0957 | -35.6629 | -37.2467
26.1018 | -35.6633 | -37.2519
26.1079 | -35.6639 | -37.2573
26.114 | -35.6644 | -37.2628
26.1201 | -35.6648 | -37.2684
26.1262 | -35.6652 | -37.2739
26.1323 | -35.6657 | -37.2793
26.1384 | -35.6661 | -37.2848
26.1445 | -35.6666 | -37.2899
26.1507 | -35.667 | -37.2957
26.1568 | -35.6675 | -37.3011
26.1629 | -35.668 | -37.3067
26.169 | -35.6685 | -37.3122
26.1751 | -35.6689 | -37.3177
26.1812 | -35.6694 | -37.3233
26.1873 | -35.6699 | -37.3284
26.1934 | -35.6704 | -37.3338
26.1995 | -35.6708 | -37.3394
26.2056 | -35.6712 | -37.3451
26.2117 | -35.6717 | -37.3506
26.2178 | -35.6722 | -37.3565
26.2239 | -35.6726 | -37.3621
26.23 | -35.673 | -37.3678
26.2361 | -35.6735 | -37.3733
26.2422 | -35.6738 | -37.3784
26.2483 | -35.6743 | -37.3835
26.2544 | -35.6748 | -37.3887
26.2605 | -35.6753 | -37.394
26.2666 | -35.6759 | -37.3993
26.2727 | -35.6763 | -37.4051
26.2788 | -35.6767 | -37.411
26.2849 | -35.6771 | -37.4167
26.2911 | -35.6776 | -37.4219
26.2972 | -35.6781 | -37.4276
26.3033 | -35.6785 | -37.433
26.3094 | -35.679 | -37.4383
26.3155 | -35.6794 | -37.4436
26.3216 | -35.6799 | -37.4488
26.3277 | -35.6802 | -37.4543
26.3338 | -35.6807 | -37.4597
26.3399 | -35.6811 | -37.4654
26.346 | -35.6815 | -37.4714
26.3521 | -35.682 | -37.477
26.3582 | -35.6825 | -37.4827
26.3643 | -35.6829 | -37.4881
26.3704 | -35.6834 | -37.4938
26.3765 | -35.6838 | -37.4988
26.3826 | -35.6843 | -37.5038
26.3887 | -35.6848 | -37.5088
26.3948 | -35.6853 | -37.5148
26.4009 | -35.6856 | -37.5211
26.407 | -35.686 | -37.5266
26.4131 | -35.6864 | -37.5327
26.4192 | -35.6869 | -37.538
26.4253 | -35.6873 | -37.5437
26.4314 | -35.6879 | -37.5497
26.4376 | -35.6883 | -37.555
26.4437 | -35.6888 | -37.5607
26.4498 | -35.6893 | -37.5661
26.4559 | -35.6897 | -37.5719
26.462 | -35.6901 | -37.5775
26.4681 | -35.6905 | -37.5829
26.4742 | -35.691 | -37.5884
26.4803 | -35.6914 | -37.5943
26.4864 | -35.6917 | -37.5992
26.4925 | -35.6922 | -37.6052
26.4986 | -35.6926 | -37.6106
26.5047 | -35.6931 | -37.6163
26.5108 | -35.6936 | -37.6218
26.5169 | -35.6941 | -37.6276
26.523 | -35.6944 | -37.6334
26.5291 | -35.6948 | -37.6389
26.5352 | -35.6953 | -37.6441
26.5413 | -35.6957 | -37.6496
26.5474 | -35.6962 | -37.6545
26.5535 | -35.6966 | -37.6604
26.5596 | -35.697 | -37.6662
26.5657 | -35.6975 | -37.6715
26.5718 | -35.6979 | -37.6774
26.578 | -35.6983 | -37.6824
26.5841 | -35.6987 | -37.6872
26.5902 | -35.6991 | -37.693
26.5963 | -35.6994 | -37.6992
26.6024 | -35.6999 | -37.7045
26.6085 | -35.7003 | -37.7105
26.6146 | -35.7007 | -37.7161
26.6207 | -35.7011 | -37.7217
26.6268 | -35.7016 | -37.7273
26.6329 | -35.702 | -37.7326
26.639 | -35.7024 | -37.7383
26.6451 | -35.7029 | -37.7446
26.6512 | -35.7034 | -37.7501
26.6573 | -35.7039 | -37.7555
26.6634 | -35.7042 | -37.7611
26.6695 | -35.7046 | -37.7668
26.6756 | -35.705 | -37.7722
26.6817 | -35.7054 | -37.7778
26.6878 | -35.7058 | -37.7838
26.6939 | -35.7062 | -37.7893
26.7 | -35.7065 | -37.7952
26.7061 | -35.707 | -37.801
26.7122 | -35.7074 | -37.8068
26.7183 | -35.7078 | -37.812
26.7245 | -35.7082 | -37.8173
26.7306 | -35.7086 | -37.8229
26.7367 | -35.709 | -37.8292
26.7428 | -35.7094 | -37.8357
26.7489 | -35.7098 | -37.8413
26.755 | -35.7101 | -37.8462
26.7611 | -35.7106 | -37.8512
26.7672 | -35.711 | -37.857
26.7733 | -35.7114 | -37.8617
26.7794 | -35.7118 | -37.8679
26.7855 | -35.7122 | -37.874
26.7916 | -35.7127 | -37.8799
26.7977 | -35.7132 | -37.8859
26.8038 | -35.7136 | -37.8916
26.8099 | -35.7139 | -37.8968
26.816 | -35.7144 | -37.903
26.8221 | -35.7148 | -37.908
26.8282 | -35.7152 | -37.9134
26.8343 | -35.7155 | -37.9185
26.8404 | -35.7158 | -37.9237
26.8465 | -35.7163 | -37.9297
26.8526 | -35.7168 | -37.9351
26.8587 | -35.7172 | -37.9409
26.8649 | -35.7175 | -37.9464
26.871 | -35.7179 | -37.9514
26.8771 | -35.7182 | -37.9567
26.8832 | -35.7186 | -37.9623
26.8893 | -35.7191 | -37.9679
26.8954 | -35.7194 | -37.9733
26.9015 | -35.7197 | -37.9796
26.9076 | -35.7201 | -37.9855
26.9137 | -35.7205 | -37.9909
26.9198 | -35.721 | -37.997
26.9259 | -35.7213 | -38.0025
26.932 | -35.7218 | -38.0082
26.9381 | -35.7222 | -38.0133
26.9442 | -35.7226 | -38.0192
26.9503 | -35.723 | -38.0242
26.9564 | -35.7234 | -38.0296
26.9625 | -35.7238 | -38.0352
26.9686 | -35.7241 | -38.0408
26.9747 | -35.7246 | -38.0459
26.9808 | -35.7248 | -38.0514
26.9869 | -35.7252 | -38.0568
26.993 | -35.7256 | -38.0627
26.9991 | -35.726 | -38.0685
27.0052 | -35.7263 | -38.074
27.0114 | -35.7267 | -38.0799
27.0175 | -35.7271 | -38.0853
27.0236 | -35.7275 | -38.0909
27.0297 | -35.7279 | -38.0968
27.0358 | -35.7283 | -38.1022
27.0419 | -35.7288 | -38.1075
27.048 | -35.7291 | -38.113
27.0541 | -35.7294 | -38.1184
27.0602 | -35.7298 | -38.124
27.0663 | -35.7303 | -38.1299
27.0724 | -35.7307 | -38.1351
27.0785 | -35.7311 | -38.1409
27.0846 | -35.7313 | -38.1469
27.0907 | -35.7317 | -38.1522
27.0968 | -35.7321 | -38.1577
27.1029 | -35.7326 | -38.1638
27.109 | -35.733 | -38.1697
27.1151 | -35.7334 | -38.176
27.1212 | -35.7337 | -38.1821
27.1273 | -35.7341 | -38.1881
27.1334 | -35.7344 | -38.1935
27.1395 | -35.7348 | -38.1991
27.1456 | -35.7352 | -38.2048
27.1518 | -35.7355 | -38.2104
27.1579 | -35.7359 | -38.2163
27.164 | -35.7363 | -38.2221
27.1701 | -35.7368 | -38.2273
27.1762 | -35.7371 | -38.2333
27.1823 | -35.7375 | -38.2387
27.1884 | -35.7378 | -38.2443
27.1945 | -35.7383 | -38.2494
27.2006 | -35.7388 | -38.255
27.2067 | -35.7392 | -38.2609
27.2128 | -35.7397 | -38.2668
27.2189 | -35.7401 | -38.2723
27.225 | -35.7404 | -38.2782
27.2311 | -35.7409 | -38.2833
27.2372 | -35.7413 | -38.2892
27.2433 | -35.7416 | -38.2946
27.2494 | -35.742 | -38.2999
27.2555 | -35.7424 | -38.3058
27.2616 | -35.7427 | -38.311
27.2677 | -35.7431 | -38.3167
27.2738 | -35.7435 | -38.3221
27.2799 | -35.7439 | -38.3277
27.286 | -35.7443 | -38.3328
27.2921 | -35.7447 | -38.3386
27.2983 | -35.745 | -38.3444
27.3044 | -35.7454 | -38.3499
27.3105 | -35.7458 | -38.3555
27.3166 | -35.7461 | -38.3612
27.3227 | -35.7464 | -38.3665
27.3288 | -35.7469 | -38.3725
27.3349 | -35.7473 | -38.3774
27.341 | -35.7477 | -38.3827
27.3471 | -35.7481 | -38.3883
27.3532 | -35.7485 | -38.3939
27.3593 | -35.7488 | -38.3993
27.3654 | -35.7492 | -38.4051
27.3715 | -35.7496 | -38.4108
27.3776 | -35.7499 | -38.4159
27.3837 | -35.7504 | -38.4208
27.3898 | -35.7507 | -38.4266
27.3959 | -35.7511 | -38.4319
27.402 | -35.7513 | -38.4369
27.4081 | -35.7517 | -38.4425
27.4142 | -35.7521 | -38.4474
27.4203 | -35.7525 | -38.4531
27.4264 | -35.7528 | -38.4585
27.4325 | -35.7532 | -38.4645
27.4387 | -35.7536 | -38.4703
27.4448 | -35.7539 | -38.4759
27.4509 | -35.7543 | -38.4814
27.457 | -35.7546 | -38.4869
27.4631 | -35.7551 | -38.4928
27.4692 | -35.7555 | -38.4983
27.4753 | -35.7558 | -38.5039
27.4814 | -35.7562 | -38.5096
27.4875 | -35.7565 | -38.5144
27.4936 | -35.7569 | -38.5203
27.4997 | -35.7572 | -38.525
27.5058 | -35.7577 | -38.5304
27.5119 | -35.7579 | -38.5355
27.518 | -35.7583 | -38.5411
27.5241 | -35.7587 | -38.5463
27.5302 | -35.7592 | -38.552
27.5363 | -35.7595 | -38.5576
27.5424 | -35.7599 | -38.5637
27.5485 | -35.7602 | -38.5697
27.5546 | -35.7606 | -38.5754
27.5607 | -35.7611 | -38.5811
27.5668 | -35.7615 | -38.5868
27.5729 | -35.7618 | -38.5923
27.5791 | -35.7622 | -38.5974
27.5852 | -35.7625 | -38.6036
27.5913 | -35.7629 | -38.6093
27.5974 | -35.7633 | -38.6147
27.6035 | -35.7636 | -38.6196
27.6096 | -35.764 | -38.6244
27.6157 | -35.7642 | -38.6297
27.6218 | -35.7645 | -38.635
27.6279 | -35.7648 | -38.6408
27.634 | -35.7651 | -38.6459
27.6401 | -35.7654 | -38.652
27.6462 | -35.7659 | -38.6574
27.6523 | -35.7662 | -38.6631
27.6584 | -35.7665 | -38.6688
27.6645 | -35.7668 | -38.6745
27.6706 | -35.7671 | -38.6804
27.6767 | -35.7675 | -38.6859
27.6828 | -35.7678 | -38.691
27.6889 | -35.7682 | -38.6958
27.695 | -35.7686 | -38.7011
27.7011 | -35.7689 | -38.7068
27.7072 | -35.7693 | -38.7123
27.7133 | -35.7697 | -38.7182
27.7194 | -35.7699 | -38.7244
27.7256 | -35.7704 | -38.7302
27.7317 | -35.7707 | -38.7353
27.7378 | -35.7711 | -38.7405
27.7439 | -35.7714 | -38.7467
27.75 | -35.7718 | -38.7523
27.7561 | -35.7721 | -38.7586
27.7622 | -35.7724 | -38.7646
27.7683 | -35.7727 | -38.7703
27.7744 | -35.7732 | -38.7759
27.7805 | -35.7735 | -38.7817
27.7866 | -35.7738 | -38.7868
27.7927 | -35.7741 | -38.7923
27.7988 | -35.7744 | -38.7981
27.8049 | -35.7748 | -38.8037
27.811 | -35.7751 | -38.8094
27.8171 | -35.7754 | -38.8149
27.8232 | -35.7758 | -38.8201
27.8293 | -35.7761 | -38.8257
27.8354 | -35.7765 | -38.8312
27.8415 | -35.7768 | -38.8363
27.8476 | -35.7772 | -38.8416
27.8537 | -35.7776 | -38.8475
27.8598 | -35.7778 | -38.8531
27.866 | -35.7782 | -38.8583
27.8721 | -35.7786 | -38.864
27.8782 | -35.779 | -38.8694
27.8843 | -35.7793 | -38.8745
27.8904 | -35.7797 | -38.8798
27.8965 | -35.78 | -38.8859
27.9026 | -35.7804 | -38.8917
27.9087 | -35.7807 | -38.8976
27.9148 | -35.781 | -38.9035
27.9209 | -35.7813 | -38.9091
27.927 | -35.7817 | -38.9147
27.9331 | -35.7819 | -38.9197
27.9392 | -35.7823 | -38.9248
27.9453 | -35.7826 | -38.9302
27.9514 | -35.7829 | -38.9361
27.9575 | -35.7833 | -38.9414
27.9636 | -35.7836 | -38.9468
27.9697 | -35.784 | -38.9521
27.9758 | -35.7844 | -38.9577
27.9819 | -35.7847 | -38.9632
27.988 | -35.785 | -38.9688
27.9941 | -35.7854 | -38.9743
28.0002 | -35.7856 | -38.9795
28.0063 | -35.7859 | -38.9851
28.0125 | -35.7862 | -38.9903
28.0186 | -35.7865 | -38.9959
28.0247 | -35.7869 | -39.0013
28.0308 | -35.7873 | -39.0067
28.0369 | -35.7875 | -39.0124
28.043 | -35.7879 | -39.0178
28.0491 | -35.7882 | -39.0237
28.0552 | -35.7886 | -39.0293
28.0613 | -35.7889 | -39.035
28.0674 | -35.7893 | -39.0403
28.0735 | -35.7896 | -39.0455

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5672 | -43.3426
21.4908 | -35.6516 | -35.7693
16.4948 | -34.9868 | -28.2519
12.6602 | -34.5931 | -21.8606
9.7171 | -34.3622 | -16.8501
7.4582 | -34.2074 | -12.7932
5.7244 | -34.0967 | -9.5136
4.3936 | -34.0272 | -6.8588
3.3722 | -34.013 | -4.7712
2.5883 | -34.0073 | -3.1898
1.9866 | -34.0026 | -1.9275
1.5248 | -34.0046 | -0.91802
1.1703 | -34.0121 | -0.25925
1 | -34.011 | 0.070005
0.89822 | -34.0042 | 0.34751
0.68941 | -33.9942 | 0.93435
0.52919 | -33.9993 | 1.3877
0.40612 | -34.0024 | 1.9929
0.31171 | -33.9928 | 2.4378
0.23928 | -33.9974 | 3.168
0.18358 | -34.0034 | 3.5774
0.14091 | -34.0134 | 4.4479
0.10819 | -34.0262 | 5.6016
0.083017 | -34.0634 | 7.0644
0.063705 | -34.0873 | 8.94
0.048923 | -34.1602 | 11.4192
0.037479 | -34.2731 | 14.5974
0.028801 | -34.443 | 18.6508
0.022078 | -34.7525 | 24.0773
0.017023 | -35.2429 | 30.4764
0.013018 | -36.0475 | 38.1399
0.010014 | -37.1599 | 46.0587

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003743 | -3.0921
0.007486 | -27.8717
0.011229 | -35.7732
0.014971 | -40.1239
0.018714 | -42.0454
0.022457 | -44.1614
0.0262 | -46.0903
0.029943 | -47.4199
0.033686 | -49.555
0.037429 | -51.3976
0.041172 | -51.0703
0.044914 | -51.7748
0.048657 | -54.0364
0.0524 | -55.0431
0.056143 | -55.5348
0.059886 | -56.1703
0.063629 | -57.137
0.067372 | -57.4288
0.071114 | -58.0677
0.074857 | -58.2516
0.0786 | -58.7587
0.082343 | -60.2008
0.086086 | -61.3657
0.089829 | -61.329
0.093572 | -61.5113
0.097314 | -61.6118
0.10106 | -62.3173
0.1048 | -63.6906
0.10854 | -63.6789
0.11229 | -63.264
0.11603 | -63.7699
0.11977 | -64.3509
0.12351 | -64.744
0.12726 | -65.1402
0.131 | -66.0235
0.13474 | -66.0595
0.13849 | -66.8597
0.14223 | -67.1644
0.14597 | -68.0213
0.14971 | -67.5025
0.15346 | -67.035
0.1572 | -66.2586
0.16094 | -67.5967
0.16469 | -67.2406
0.16843 | -68.1652
0.17217 | -69.4972
0.17591 | -68.8641
0.17966 | -69.1694
0.1834 | -69.728
0.18714 | -69.1894
0.19089 | -70.0475
0.19463 | -70.6621
0.19837 | -70.3154
0.20211 | -70.5405
0.20586 | -70.1263
0.2096 | -70.4734
0.21334 | -70.4513
0.21709 | -70.7299
0.22083 | -71.9118
0.22457 | -72.2022
0.22831 | -71.7805
0.23206 | -71.3263
0.2358 | -71.6298
0.23954 | -71.9723
0.24329 | -72.2987
0.24703 | -73.4665
0.25077 | -73.1301
0.25451 | -72.4909
0.25826 | -73.1293
0.262 | -72.9654
0.26574 | -73.2776
0.26949 | -73.2796
0.27323 | -73.1135
0.27697 | -73.3749
0.28071 | -74.287
0.28446 | -74.0612
0.2882 | -74.865
0.29194 | -74.2251
0.29569 | -73.9727
0.29943 | -75.6374
0.30317 | -76.2574
0.30691 | -75.2246
0.31066 | -74.9653
0.3144 | -75.6693
0.31814 | -75.7327
0.32189 | -76.7685
0.32563 | -75.918
0.32937 | -75.5517
0.33311 | -75.0487
0.33686 | -75.9288
0.3406 | -76.5205
0.34434 | -76.4257
0.34809 | -77.1197
0.35183 | -76.762
0.35557 | -75.5774
0.35931 | -75.9322
0.36306 | -76.0583
0.3668 | -75.6332
0.37054 | -75.9215
0.37429 | -76.5296
0.37803 | -76.1454
0.38177 | -76.0462
0.38551 | -75.7069
0.38926 | -75.9312
0.393 | -76.7284
0.39674 | -76.327
0.40049 | -77.3022
0.40423 | -77.2503
0.40797 | -77.3639
0.41171 | -77.664
0.41546 | -77.3065
0.4192 | -77.4359
0.42294 | -78.6125
0.42669 | -79.0343
0.43043 | -78.0483
0.43417 | -77.9243
0.43791 | -77.8482
0.44166 | -78.0904
0.4454 | -78.6018
0.44914 | -79.2985
0.45289 | -79.8388
0.45663 | -79.6756
0.46037 | -78.7588
0.46411 | -79.0906
0.46786 | -79.8522
0.4716 | -79.187
0.47534 | -79.7845
0.47909 | -79.5626
0.48283 | -79.1598
0.48657 | -79.0135
0.49032 | -78.8079
0.49406 | -79.6284
0.4978 | -80.6874
0.50154 | -80.8586
0.50529 | -80.1674
0.50903 | -80.6236
0.51277 | -80.9365
0.51651 | -78.9727
0.52026 | -79.0577
0.524 | -80.8878
0.52774 | -80.3595
0.53149 | -80.5439
0.53523 | -80.7572
0.53897 | -81.4287
0.54271 | -82.1347
0.54646 | -81.7853
0.5502 | -81.0709
0.55394 | -80.8188
0.55769 | -81.1049
0.56143 | -81.3437
0.56517 | -81.4061
0.56892 | -81.3412
0.57266 | -81.1949
0.5764 | -81.2283
0.58014 | -81.3863
0.58389 | -81.3973
0.58763 | -81.174
0.59137 | -81.0884
0.59512 | -81.5003
0.59886 | -81.8435
0.6026 | -81.3607
0.60634 | -80.9704
0.61009 | -81.0672
0.61383 | -80.989
0.61757 | -80.8089
0.62132 | -81.0674
0.62506 | -81.117
0.6288 | -81.2113
0.63254 | -81.4508
0.63629 | -81.612
0.64003 | -81.9054
0.64377 | -82.4242
0.64752 | -82.6678
0.65126 | -82.2394
0.655 | -81.8288
0.65874 | -81.7833
0.66249 | -81.539
0.66623 | -81.7001
0.66997 | -82.8896
0.67372 | -84.1791
0.67746 | -84.1189
0.6812 | -83.2551
0.68494 | -83.1058
0.68869 | -83.5719
0.69243 | -83.7387
0.69617 | -83.5447
0.69992 | -83.3535
0.70366 | -83.4799
0.7074 | -83.2406
0.71114 | -82.4013
0.71489 | -81.9162
0.71863 | -82.1999
0.72237 | -82.8173
0.72612 | -83.3589
0.72986 | -83.6615
0.7336 | -82.9988
0.73734 | -82.5145
0.74109 | -82.8296
0.74483 | -83.2106
0.74857 | -83.3721
0.75232 | -83.4433
0.75606 | -83.6357
0.7598 | -83.7778
0.76354 | -83.6721
0.76729 | -83.6134
0.77103 | -84.0433
0.77477 | -84.4513
0.77852 | -84.4048
0.78226 | -84.3161
0.786 | -84.3734
0.78975 | -84.6731
0.79349 | -84.7119
0.79723 | -84.1477
0.80097 | -83.783
0.80472 | -84.0451
0.80846 | -84.9631
0.8122 | -85.4736
0.81595 | -85.1486
0.81969 | -84.7949
0.82343 | -84.4244
0.82717 | -84.1287
0.83092 | -84.1995
0.83466 | -84.5695
0.8384 | -84.858
0.84215 | -84.6988
0.84589 | -84.1696
0.84963 | -84.2217
0.85337 | -84.8116
0.85712 | -85.3169
0.86086 | -85.3347
0.8646 | -85.028
0.86835 | -84.8931
0.87209 | -84.7484
0.87583 | -84.5969
0.87957 | -84.5031
0.88332 | -84.4896
0.88706 | -84.826
0.8908 | -85.0571
0.89455 | -84.9686
0.89829 | -85.1329
0.90203 | -85.4686
0.90577 | -85.2701
0.90952 | -85.1752
0.91326 | -84.8859
0.917 | -84.6324
0.92075 | -84.7532
0.92449 | -85.1727
0.92823 | -85.68
0.93197 | -86.1635
0.93572 | -86.5122
0.93946 | -86.4975
0.9432 | -86.3141
0.94695 | -86.0272
0.95069 | -86.1443
0.95443 | -86.3322
0.95817 | -86.1925
0.96192 | -86.3427
0.96566 | -86.1852
0.9694 | -85.6883
0.97315 | -85.6558
0.97689 | -86.027
0.98063 | -86.4624
0.98437 | -86.764
0.98812 | -86.6994
0.99186 | -86.4313
0.9956 | -86.294
0.99935 | -86.4388
1.0031 | -87.0766
1.0068 | -87.5721
1.0106 | -87.324
1.0143 | -86.862
1.0181 | -86.675
1.0218 | -86.8655
1.0255 | -87.2707
1.0293 | -87.5361
1.033 | -87.4999
1.0368 | -86.6643
1.0405 | -86.073
1.0443 | -86.2819
1.048 | -86.8447
1.0517 | -87.2644
1.0555 | -87.4196
1.0592 | -87.0941
1.063 | -86.807
1.0667 | -87.3526
1.0705 | -87.9592
1.0742 | -87.6461
1.0779 | -87.1099
1.0817 | -86.7702
1.0854 | -86.6744
1.0892 | -86.9092
1.0929 | -87.2972
1.0967 | -87.5248
1.1004 | -87.5854
1.1041 | -87.5517
1.1079 | -87.3661
1.1116 | -87.0617
1.1154 | -86.8523
1.1191 | -86.8948
1.1229 | -87.1933
1.1266 | -87.6449
1.1303 | -87.9312
1.1341 | -87.7828
1.1378 | -87.5018
1.1416 | -87.5014
1.1453 | -87.8557
1.1491 | -88.2928
1.1528 | -88.4579
1.1565 | -88.2388
1.1603 | -87.897
1.164 | -87.7818
1.1678 | -87.9674
1.1715 | -88.2132
1.1753 | -88.3257
1.179 | -88.3617
1.1827 | -88.4002
1.1865 | -88.3815
1.1902 | -88.2639
1.194 | -88.1663
1.1977 | -88.211
1.2015 | -88.2362
1.2052 | -88.0364
1.2089 | -87.7421
1.2127 | -87.6454
1.2164 | -87.8893
1.2202 | -88.3542
1.2239 | -88.7246
1.2277 | -88.8072
1.2314 | -88.6962
1.2351 | -88.5852
1.2389 | -88.6359
1.2426 | -88.8352
1.2464 | -88.9456
1.2501 | -88.8005
1.2539 | -88.5301
1.2576 | -88.4019
1.2613 | -88.5001
1.2651 | -88.6453
1.2688 | -88.7022
1.2726 | -88.7164
1.2763 | -88.7821
1.2801 | -89.0251
1.2838 | -89.4533
1.2875 | -89.8583
1.2913 | -89.9765
1.295 | -89.7633
1.2988 | -89.392
1.3025 | -89.0093
1.3063 | -88.7257
1.31 | -88.5803
1.3137 | -88.4717
1.3175 | -88.3889
1.3212 | -88.4448
1.325 | -88.6451
1.3287 | -88.8918
1.3325 | -89.0824
1.3362 | -89.198
1.3399 | -89.3048
1.3437 | -89.4709
1.3474 | -89.5583
1.3512 | -89.328
1.3549 | -88.9961
1.3587 | -88.8949
1.3624 | -88.92
1.3661 | -88.9281
1.3699 | -88.9804
1.3736 | -89.044
1.3774 | -89.033
1.3811 | -89.0266
1.3849 | -89.1291
1.3886 | -89.3426
1.3923 | -89.5134
1.3961 | -89.4318
1.3998 | -89.0751
1.4036 | -88.6822
1.4073 | -88.5594
1.4111 | -88.8555
1.4148 | -89.4171
1.4185 | -89.9629
1.4223 | -90.2951
1.426 | -90.2442
1.4298 | -89.8961
1.4335 | -89.5998
1.4373 | -89.5473
1.441 | -89.7493
1.4447 | -90.1283
1.4485 | -90.4789
1.4522 | -90.5092
1.456 | -90.2124
1.4597 | -89.9023
1.4635 | -89.8016
1.4672 | -89.9744
1.4709 | -90.2828
1.4747 | -90.454
1.4784 | -90.3819
1.4822 | -90.2249
1.4859 | -90.2143
1.4897 | -90.3604
1.4934 | -90.4066
1.4971 | -90.1955
1.5009 | -89.9379
1.5046 | -89.8724
1.5084 | -89.9794
1.5121 | -89.9964
1.5159 | -89.7495
1.5196 | -89.464
1.5233 | -89.4889
1.5271 | -89.8884
1.5308 | -90.4013
1.5346 | -90.7161
1.5383 | -90.8739
1.5421 | -91.0026
1.5458 | -91.1079
1.5495 | -91.2137
1.5533 | -91.2461
1.557 | -91.1255
1.5608 | -90.9947
1.5645 | -90.9515
1.5683 | -90.9055
1.572 | -90.8004
1.5757 | -90.7235
1.5795 | -90.7241
1.5832 | -90.7498
1.587 | -90.7208
1.5907 | -90.6507
1.5945 | -90.7153
1.5982 | -90.9583
1.6019 | -91.1643
1.6057 | -91.1785
1.6094 | -91.1086
1.6132 | -91.067
1.6169 | -91.1087
1.6207 | -91.2076
1.6244 | -91.3622
1.6281 | -91.4351
1.6319 | -91.3328
1.6356 | -91.1435
1.6394 | -91.005
1.6431 | -90.9756
1.6469 | -91.0501
1.6506 | -91.1996
1.6543 | -91.3768
1.6581 | -91.5159
1.6618 | -91.5716
1.6656 | -91.5206
1.6693 | -91.3602
1.6731 | -91.106
1.6768 | -90.8465
1.6805 | -90.6966
1.6843 | -90.73
1.688 | -90.9621
1.6918 | -91.3505
1.6955 | -91.7726
1.6993 | -92.0517
1.703 | -92.1174
1.7067 | -92.0759
1.7105 | -92.0638
1.7142 | -92.1075
1.718 | -92.1058
1.7217 | -91.9737
1.7255 | -91.7157
1.7292 | -91.4286
1.7329 | -91.228
1.7367 | -91.1762
1.7404 | -91.2635
1.7442 | -91.429
1.7479 | -91.5696
1.7517 | -91.5811
1.7554 | -91.4281
1.7591 | -91.1864
1.7629 | -90.9678
1.7666 | -90.8101
1.7704 | -90.6788
1.7741 | -90.6023
1.7779 | -90.6631
1.7816 | -90.9078
1.7853 | -91.29
1.7891 | -91.6958
1.7928 | -92.0192
1.7966 | -92.2056
1.8003 | -92.33
1.8041 | -92.4522
1.8078 | -92.504
1.8115 | -92.4346
1.8153 | -92.2441
1.819 | -91.9426
1.8228 | -91.5895
1.8265 | -91.3394
1.8303 | -91.3339
1.834 | -91.5441
1.8377 | -91.8296
1.8415 | -92.0838
1.8452 | -92.2817
1.849 | -92.3806
1.8527 | -92.2978
1.8565 | -92.0939
1.8602 | -91.979
1.8639 | -92.1044
1.8677 | -92.4428
1.8714 | -92.8322
1.8752 | -93.0782
1.8789 | -93.0952
1.8827 | -92.9315
1.8864 | -92.6783
1.8901 | -92.3924
1.8939 | -92.1242
1.8976 | -91.9468
1.9014 | -91.8676
1.9051 | -91.8391
1.9089 | -91.8179
1.9126 | -91.7804
1.9163 | -91.7589
1.9201 | -91.8357
1.9238 | -92.0728
1.9276 | -92.4232
1.9313 | -92.7013
1.9351 | -92.7962
1.9388 | -92.7745
1.9425 | -92.7443
1.9463 | -92.7223
1.95 | -92.7115
1.9538 | -92.7313
1.9575 | -92.7951
1.9613 | -92.8859
1.965 | -92.9657
1.9687 | -92.9778
1.9725 | -92.8829
1.9762 | -92.717
1.98 | -92.5391
1.9837 | -92.3746
1.9875 | -92.2659
1.9912 | -92.2428
1.9949 | -92.322
1.9987 | -92.4725
2.0024 | -92.5987
2.0062 | -92.6339
2.0099 | -92.6172
2.0137 | -92.6297
2.0174 | -92.6775
2.0211 | -92.6879
2.0249 | -92.635
2.0286 | -92.5309
2.0324 | -92.3828
2.0361 | -92.2504
2.0399 | -92.2121
2.0436 | -92.3183
2.0473 | -92.5445
2.0511 | -92.7842
2.0548 | -92.9126
2.0586 | -92.8955
2.0623 | -92.8176
2.0661 | -92.7762
2.0698 | -92.7729
2.0735 | -92.7889
2.0773 | -92.8365
2.081 | -92.9578
2.0848 | -93.1763
2.0885 | -93.4544
2.0923 | -93.7167
2.096 | -93.9118
2.0997 | -93.986
2.1035 | -93.8932
2.1072 | -93.6515
2.111 | -93.3763
2.1147 | -93.2138
2.1185 | -93.2301
2.1222 | -93.372
2.1259 | -93.5602
2.1297 | -93.7603
2.1334 | -93.954
2.1372 | -94.0775
2.1409 | -94.0773
2.1447 | -93.9521
2.1484 | -93.7585
2.1521 | -93.532
2.1559 | -93.3598
2.1596 | -93.3876
2.1634 | -93.4373
2.1671 | -93.5701
2.1709 | -93.7437
2.1746 | -93.8952
2.1783 | -93.9721
2.1821 | -93.9532
2.1858 | -93.848
2.1896 | -93.7069
2.1933 | -93.5623
2.1971 | -93.4453
2.2008 | -93.3547
2.2045 | -93.2594
2.2083 | -93.1604
2.212 | -93.0717
2.2158 | -92.9961
2.2195 | -92.9463
2.2233 | -92.9587
2.227 | -93.0509
2.2307 | -93.214
2.2345 | -93.4046
2.2382 | -93.5807
2.242 | -93.7138
2.2457 | -93.7851
2.2495 | -93.7997
2.2532 | -93.7906
2.2569 | -93.827
2.2607 | -93.9309
2.2644 | -94.0768
2.2682 | -94.2266
2.2719 | -94.3305
2.2757 | -94.3759
2.2794 | -94.3771
2.2831 | -94.3648
2.2869 | -94.3397
2.2906 | -94.2658
2.2944 | -94.161
2.2981 | -94.0573
2.3019 | -94.0119
2.3056 | -94.031
2.3093 | -94.0899
2.3131 | -94.1513
2.3168 | -94.1906
2.3206 | -94.1921
2.3243 | -94.1609
2.3281 | -94.0926
2.3318 | -94
2.3355 | -93.8797
2.3393 | -93.7661
2.343 | -93.6961
2.3468 | -93.7138
2.3505 | -93.8245
2.3543 | -93.9993
2.358 | -94.1442
2.3617 | -94.1862
2.3655 | -94.0963
2.3692 | -93.9545
2.373 | -93.8451
2.3767 | -93.8176
2.3805 | -93.8373
2.3842 | -93.8824
2.3879 | -93.9318
2.3917 | -93.9879
2.3954 | -94.0399
2.3992 | -94.0743
2.4029 | -94.0865
2.4067 | -94.097
2.4104 | -94.1219
2.4141 | -94.174
2.4179 | -94.2302
2.4216 | -94.2581
2.4254 | -94.24
2.4291 | -94.2054
2.4329 | -94.194
2.4366 | -94.2328
2.4403 | -94.3039
2.4441 | -94.3437
2.4478 | -94.308
2.4516 | -94.1903
2.4553 | -94.0283
2.4591 | -93.89
2.4628 | -93.8239
2.4665 | -93.8173
2.4703 | -93.8665
2.474 | -93.9466
2.4778 | -94.0287
2.4815 | -94.0663
2.4853 | -94.05
2.489 | -94.0064
2.4927 | -93.9975
2.4965 | -94.0526
2.5002 | -94.1508
2.504 | -94.2345
2.5077 | -94.2779
2.5115 | -94.2785
2.5152 | -94.2346
2.5189 | -94.1417
2.5227 | -94.0101
2.5264 | -93.861
2.5302 | -93.7541
2.5339 | -93.7442
2.5377 | -93.8635
2.5414 | -94.0733
2.5451 | -94.3047
2.5489 | -94.4872
2.5526 | -94.6141
2.5564 | -94.7332
2.5601 | -94.8615
2.5639 | -94.9441
2.5676 | -94.9201
2.5713 | -94.7977
2.5751 | -94.6586
2.5788 | -94.5736
2.5826 | -94.5531
2.5863 | -94.5612
2.5901 | -94.5742
2.5938 | -94.5568
2.5975 | -94.4692
2.6013 | -94.2988
2.605 | -94.0865
2.6088 | -93.9217
2.6125 | -93.8781
2.6163 | -93.974
2.62 | -94.1902
2.6237 | -94.4752
2.6275 | -94.7661
2.6312 | -95.0031
2.635 | -95.1656
2.6387 | -95.2405
2.6425 | -95.218
2.6462 | -95.1099
2.6499 | -94.9029
2.6537 | -94.6247
2.6574 | -94.3286
2.6612 | -94.077
2.6649 | -93.9514
2.6687 | -93.9717
2.6724 | -94.1079
2.6761 | -94.3019
2.6799 | -94.4944
2.6836 | -94.6391
2.6874 | -94.7089
2.6911 | -94.7431
2.6949 | -94.7967
2.6986 | -94.915
2.7023 | -95.004
2.7061 | -95.0599
2.7098 | -95.0692
2.7136 | -95.0473
2.7173 | -95.029
2.7211 | -95.0265
2.7248 | -95.0464
2.7285 | -95.0823
2.7323 | -95.1319
2.736 | -95.169
2.7398 | -95.1846
2.7435 | -95.1669
2.7473 | -95.1153
2.751 | -95.0672
2.7547 | -95.0318
2.7585 | -94.9943
2.7622 | -94.9496
2.766 | -94.8919
2.7697 | -94.8191
2.7735 | -94.7312
2.7772 | -94.6431
2.7809 | -94.5824
2.7847 | -94.5586
2.7884 | -94.5678
2.7922 | -94.6133
2.7959 | -94.6874
2.7997 | -94.7762
2.8034 | -94.9005
2.8071 | -95.0464
2.8109 | -95.2153
2.8146 | -95.4068
2.8184 | -95.6353
2.8221 | -95.9043
2.8259 | -96.1655
2.8296 | -96.3457
2.8333 | -96.416
2.8371 | -96.3703
2.8408 | -96.2447
2.8446 | -96.1143
2.8483 | -96.0166
2.8521 | -95.971
2.8558 | -95.9909
2.8595 | -96.0371
2.8633 | -96.0842
2.867 | -96.1066
2.8708 | -96.0945
2.8745 | -96.0562
2.8783 | -95.9944
2.882 | -95.8974
2.8857 | -95.752
2.8895 | -95.5694
2.8932 | -95.3806
2.897 | -95.2341
2.9007 | -95.1513
2.9045 | -95.1387
2.9082 | -95.1999
2.9119 | -95.3134
2.9157 | -95.4576
2.9194 | -95.5966
2.9232 | -95.6966
2.9269 | -95.7297
2.9307 | -95.7251
2.9344 | -95.7279
2.9381 | -95.7401
2.9419 | -95.7594
2.9456 | -95.7874
2.9494 | -95.8221
2.9531 | -95.8633
2.9569 | -95.9031
2.9606 | -95.9328
2.9643 | -95.933
2.9681 | -95.8886
2.9718 | -95.8093
2.9756 | -95.7093
2.9793 | -95.6071
2.9831 | -95.5355
2.9868 | -95.528
2.9905 | -95.5736
2.9943 | -95.6562
2.998 | -95.7437
3.0018 | -95.8067
3.0055 | -95.8246
3.0093 | -95.8082
3.013 | -95.7864
3.0167 | -95.7691
3.0205 | -95.7684
3.0242 | -95.7841
3.028 | -95.7934
3.0317 | -95.7748
3.0355 | -95.7238
3.0392 | -95.6667
3.0429 | -95.6317
3.0467 | -95.6417
3.0504 | -95.6914
3.0542 | -95.7826
3.0579 | -95.906
3.0617 | -96.0475
3.0654 | -96.1953
3.0691 | -96.2983
3.0729 | -96.3501
3.0766 | -96.3791
3.0804 | -96.4201
3.0841 | -96.4911
3.0879 | -96.5812
3.0916 | -96.6599
3.0953 | -96.6903
3.0991 | -96.6494
3.1028 | -96.5418
3.1066 | -96.4234
3.1103 | -96.3482
3.1141 | -96.3183
3.1178 | -96.3151
3.1216 | -96.316
3.1253 | -96.3113
3.129 | -96.323
3.1328 | -96.3548
3.1365 | -96.4087
3.1403 | -96.4746
3.144 | -96.5669
3.1477 | -96.6903
3.1515 | -96.8218
3.1552 | -96.8988
3.159 | -96.8879
3.1627 | -96.8029
3.1665 | -96.6867
3.1702 | -96.583
3.174 | -96.5095
3.1777 | -96.4692
3.1814 | -96.4679
3.1852 | -96.496
3.1889 | -96.5588
3.1927 | -96.6534
3.1964 | -96.7702
3.2001 | -96.8969
3.2039 | -97.0101
3.2076 | -97.0831
3.2114 | -97.1041
3.2151 | -97.0623
3.2189 | -96.947
3.2226 | -96.802
3.2264 | -96.6563
3.2301 | -96.5547
3.2338 | -96.5236
3.2376 | -96.621
3.2413 | -96.6911
3.2451 | -96.7758
3.2488 | -96.8665
3.2525 | -96.9257
3.2563 | -96.946
3.26 | -96.9165
3.2638 | -96.85
3.2675 | -96.7646
3.2713 | -96.6862
3.275 | -96.6275
3.2788 | -96.5953
3.2825 | -96.5756
3.2862 | -96.5641
3.29 | -96.5524
3.2937 | -96.5555
3.2975 | -96.5878
3.3012 | -96.6671
3.3049 | -96.7827
3.3087 | -96.9296
3.3124 | -97.067
3.3162 | -97.1633
3.3199 | -97.1969
3.3237 | -97.1684
3.3274 | -97.0933
3.3312 | -96.977
3.3349 | -96.8555
3.3386 | -96.7551
3.3424 | -96.6845
3.3461 | -96.6416
3.3499 | -96.6155
3.3536 | -96.6029
3.3573 | -96.6058
3.3611 | -96.6336
3.3648 | -96.6904
3.3686 | -96.7777
3.3723 | -96.8837
3.3761 | -96.9957
3.3798 | -97.0882
3.3836 | -97.1347
3.3873 | -97.1481
3.391 | -97.1146
3.3948 | -97.0482
3.3985 | -96.979
3.4023 | -96.9462
3.406 | -96.9655
3.4097 | -97.0382
3.4135 | -97.1422
3.4172 | -97.2564
3.421 | -97.342
3.4247 | -97.3599
3.4285 | -97.311
3.4322 | -97.2093
3.436 | -97.1004
3.4397 | -97.0195
3.4434 | -96.9774
3.4472 | -96.9796
3.4509 | -97.0096
3.4547 | -97.0618
3.4584 | -97.1073
3.4621 | -97.1418
3.4659 | -97.1643
3.4696 | -97.1646
3.4734 | -97.1299
3.4771 | -97.0782
3.4809 | -97.0324
3.4846 | -97.0067
3.4884 | -97.0077
3.4921 | -97.047
3.4958 | -97.1324
3.4996 | -97.2612
3.5033 | -97.4368
3.5071 | -97.6374
3.5108 | -97.822
3.5145 | -97.9849
3.5183 | -98.0906
3.522 | -98.1173
3.5258 | -98.0677
3.5295 | -97.9668
3.5333 | -97.8511
3.537 | -97.7365
3.5408 | -97.6373
3.5445 | -97.5811
3.5482 | -97.5594
3.552 | -97.5581
3.5557 | -97.5867
3.5595 | -97.6366
3.5632 | -97.707
3.567 | -97.7866
3.5707 | -97.8548
3.5744 | -97.917
3.5782 | -97.9685
3.5819 | -98.0012
3.5857 | -98.0022
3.5894 | -97.9656
3.5932 | -97.9164
3.5969 | -97.8631
3.6006 | -97.806
3.6044 | -97.7736
3.6081 | -97.7674
3.6119 | -97.7772
3.6156 | -97.8004
3.6194 | -97.8173
3.6231 | -97.8097
3.6268 | -97.7682
3.6306 | -97.6871
3.6343 | -97.5976
3.6381 | -97.5162
3.6418 | -97.4518
3.6456 | -97.401
3.6493 | -97.3531
3.653 | -97.3109
3.6568 | -97.2627
3.6605 | -97.2031
3.6643 | -97.139
3.668 | -97.0814
3.6718 | -97.0476
3.6755 | -97.0576
3.6792 | -97.1199
3.683 | -97.2302
3.6867 | -97.3744
3.6905 | -97.525
3.6942 | -97.6721
3.698 | -97.7904
3.7017 | -97.8689
3.7054 | -97.8965
3.7092 | -97.8793
3.7129 | -97.8372
3.7167 | -97.8061
3.7204 | -97.8026
3.7242 | -97.8309
3.7279 | -97.8883
3.7316 | -97.9618
3.7354 | -98.0407
3.7391 | -98.1028
3.7429 | -98.1419
3.7466 | -98.1535
3.7504 | -98.1263
3.7541 | -98.0977
3.7578 | -98.0604
3.7616 | -98.0198
3.7653 | -97.968
3.7691 | -97.9096
3.7728 | -97.8294
3.7766 | -97.7636
3.7803 | -97.6971
3.784 | -97.6526
3.7878 | -97.6349
3.7915 | -97.6327
3.7953 | -97.6463
3.799 | -97.6815
3.8028 | -97.7331
3.8065 | -97.8
3.8102 | -97.8867
3.814 | -97.9655
3.8177 | -98.0531
3.8215 | -98.1288
3.8252 | -98.1819
3.829 | -98.2027
3.8327 | -98.1905
3.8364 | -98.1492
3.8402 | -98.1007
3.8439 | -98.0744
3.8477 | -98.068
3.8514 | -98.1035
3.8552 | -98.1987
3.8589 | -98.3238
3.8626 | -98.4549
3.8664 | -98.5714
3.8701 | -98.6547
3.8739 | -98.677
3.8776 | -98.6499
3.8814 | -98.574
3.8851 | -98.4707
3.8888 | -98.3541
3.8926 | -98.2265
3.8963 | -98.1009
3.9001 | -97.9845
3.9038 | -97.883
3.9076 | -97.7932
3.9113 | -97.7313
3.915 | -97.6817
3.9188 | -97.6461
3.9225 | -97.6404
3.9263 | -97.6699
3.93 | -97.7303
3.9338 | -97.8221
3.9375 | -97.953
3.9412 | -98.1037
3.945 | -98.2572
3.9487 | -98.418
3.9525 | -98.5621
3.9562 | -98.6799
3.96 | -98.7552
3.9637 | -98.7781
3.9674 | -98.7444
3.9712 | -98.6717
3.9749 | -98.5749
3.9787 | -98.4849
3.9824 | -98.4102
3.9862 | -98.3498
3.9899 | -98.3119
3.9936 | -98.297
3.9974 | -98.2983
4.0011 | -98.3145
4.0049 | -98.3469
4.0086 | -98.3966
4.0124 | -98.4442
4.0161 | -98.4965
4.0198 | -98.5496
4.0236 | -98.5984
4.0273 | -98.6409
4.0311 | -98.6876
4.0348 | -98.7316
4.0386 | -98.7739
4.0423 | -98.8203
4.046 | -98.8596
4.0498 | -98.8968
4.0535 | -98.923
4.0573 | -98.9273
4.061 | -98.9134
4.0648 | -98.8684
4.0685 | -98.794
4.0722 | -98.7132
4.076 | -98.6381
4.0797 | -98.5832
4.0835 | -98.5599
4.0872 | -98.5653
4.091 | -98.5899
4.0947 | -98.6214
4.0984 | -98.6631
4.1022 | -98.7032
4.1059 | -98.7498
4.1097 | -98.8017
4.1134 | -98.8522
4.1172 | -98.8995
4.1209 | -98.9366
4.1246 | -98.9562
4.1284 | -98.9779
4.1321 | -99.0075
4.1359 | -99.0278
4.1396 | -99.0525
4.1434 | -99.0786
4.1471 | -99.1112
4.1508 | -99.1549
4.1546 | -99.2099
4.1583 | -99.2677
4.1621 | -99.3102
4.1658 | -99.3397
4.1696 | -99.343
4.1733 | -99.3317
4.177 | -99.3133
4.1808 | -99.2938
4.1845 | -99.2707
4.1883 | -99.2249
4.192 | -99.1472
4.1958 | -99.0287
4.1995 | -98.8933
4.2032 | -98.7533
4.207 | -98.6376
4.2107 | -98.5509
4.2145 | -98.5072
4.2182 | -98.4923
4.222 | -98.4939
4.2257 | -98.5185
4.2294 | -98.5602
4.2332 | -98.6067
4.2369 | -98.6577
4.2407 | -98.7151
4.2444 | -98.7671
4.2482 | -98.8109
4.2519 | -98.8444
4.2556 | -98.8506
4.2594 | -98.8396
4.2631 | -98.8126
4.2669 | -98.7874
4.2706 | -98.7639
4.2744 | -98.7643
4.2781 | -98.7882
4.2818 | -98.8429
4.2856 | -98.9206
4.2893 | -98.9971
4.2931 | -99.0665
4.2968 | -99.1182
4.3006 | -99.1398
4.3043 | -99.1253
4.308 | -99.0819
4.3118 | -99.0209
4.3155 | -98.946
4.3193 | -98.8937
4.323 | -98.8532
4.3268 | -98.834
4.3305 | -98.8405
4.3342 | -98.8788
4.338 | -98.9532
4.3417 | -99.0492
4.3455 | -99.1539
4.3492 | -99.2549
4.353 | -99.3508
4.3567 | -99.4381
4.3604 | -99.5116
4.3642 | -99.5734
4.3679 | -99.6158
4.3717 | -99.6404
4.3754 | -99.6457
4.3792 | -99.6303
4.3829 | -99.593
4.3866 | -99.5362
4.3904 | -99.4597
4.3941 | -99.3779
4.3979 | -99.2899
4.4016 | -99.2075
4.4054 | -99.1342
4.4091 | -99.0771
4.4128 | -99.046
4.4166 | -99.0301
4.4203 | -99.0247
4.4241 | -99.0235
4.4278 | -99.021
4.4316 | -99.0252
4.4353 | -99.033
4.439 | -99.0403
4.4428 | -99.0449
4.4465 | -99.0425
4.4503 | -99.0394
4.454 | -99.0358
4.4578 | -99.0231
4.4615 | -99.0019
4.4652 | -98.9774
4.469 | -98.9594
4.4727 | -98.946
4.4765 | -98.9418
4.4802 | -98.9492
4.484 | -98.9594
4.4877 | -98.9837
4.4914 | -99.0148
4.4952 | -99.0522
4.4989 | -99.1014
4.5027 | -99.1441
4.5064 | -99.187
4.5102 | -99.2276
4.5139 | -99.2642
4.5176 | -99.3012
4.5214 | -99.3251
4.5251 | -99.3422
4.5289 | -99.3516
4.5326 | -99.3608
4.5364 | -99.3633
4.5401 | -99.3565
4.5438 | -99.331
4.5476 | -99.2973
4.5513 | -99.2538
4.5551 | -99.2146
4.5588 | -99.185
4.5626 | -99.1781
4.5663 | -99.1978
4.57 | -99.2353
4.5738 | -99.2866
4.5775 | -99.3368
4.5813 | -99.3774
4.585 | -99.3959
4.5888 | -99.3976
4.5925 | -99.3842
4.5962 | -99.3648
4.6 | -99.3403
4.6037 | -99.3145
4.6075 | -99.2946
4.6112 | -99.2894
4.615 | -99.2963
4.6187 | -99.3161
4.6224 | -99.3536
4.6262 | -99.4041
4.6299 | -99.4631
4.6337 | -99.5171
4.6374 | -99.5619
4.6412 | -99.591
4.6449 | -99.604
4.6486 | -99.607
4.6524 | -99.6093
4.6561 | -99.6197
4.6599 | -99.6354
4.6636 | -99.6607
4.6674 | -99.689
4.6711 | -99.7189
4.6748 | -99.7202
4.6786 | -99.7009
4.6823 | -99.6611
4.6861 | -99.6019
4.6898 | -99.5158
4.6936 | -99.4296
4.6973 | -99.3588
4.701 | -99.3211
4.7048 | -99.325
4.7085 | -99.3714
4.7123 | -99.4616
4.716 | -99.5788
4.7198 | -99.7062
4.7235 | -99.8233
4.7272 | -99.9215
4.731 | -99.9965
4.7347 | -100.0368
4.7385 | -100.0456
4.7422 | -100.0367
4.746 | -100.0268
4.7497 | -99.9999
4.7534 | -99.962
4.7572 | -99.9358
4.7609 | -99.9149
4.7647 | -99.893
4.7684 | -99.8699
4.7722 | -99.83
4.7759 | -99.7811
4.7796 | -99.7306
4.7834 | -99.6838
4.7871 | -99.6506
4.7909 | -99.6378
4.7946 | -99.6607
4.7984 | -99.7079
4.8021 | -99.767
4.8058 | -99.8413
4.8096 | -99.9077
4.8133 | -99.9591
4.8171 | -100
4.8208 | -100.0168
4.8246 | -100.0031
4.8283 | -99.9568
4.832 | -99.8989
4.8358 | -99.8293
4.8395 | -99.7589
4.8433 | -99.7063
4.847 | -99.6726
4.8508 | -99.6656
4.8545 | -99.6833
4.8582 | -99.7336
4.862 | -99.7676
4.8657 | -99.806
4.8695 | -99.8349
4.8732 | -99.8585
4.877 | -99.8772
4.8807 | -99.8863
4.8844 | -99.8889
4.8882 | -99.8782
4.8919 | -99.8532
4.8957 | -99.8339
4.8994 | -99.8187
4.9032 | -99.8025
4.9069 | -99.7904
4.9106 | -99.7896
4.9144 | -99.7948
4.9181 | -99.8018
4.9219 | -99.8054
4.9256 | -99.8065
4.9294 | -99.8177
4.9331 | -99.845
4.9368 | -99.8805
4.9406 | -99.934
4.9443 | -100.0048
4.9481 | -100.09
4.9518 | -100.1776
4.9556 | -100.254
4.9593 | -100.3239
4.963 | -100.3917
4.9668 | -100.4448
4.9705 | -100.4843
4.9743 | -100.5033
4.978 | -100.4988
4.9818 | -100.4772
4.9855 | -100.4316
4.9892 | -100.3712
4.993 | -100.3057
4.9967 | -100.2371
5.0005 | -100.1697
5.0042 | -100.1092
5.008 | -100.0515
5.0117 | -99.9975
5.0154 | -99.9571
5.0192 | -99.9194
5.0229 | -99.889
5.0267 | -99.8766
5.0304 | -99.871
5.0342 | -99.873
5.0379 | -99.8794
5.0416 | -99.8826
5.0454 | -99.8933
5.0491 | -99.9047
5.0529 | -99.9184
5.0566 | -99.9391
5.0604 | -99.9664
5.0641 | -99.9993
5.0678 | -100.0366
5.0716 | -100.076
5.0753 | -100.121
5.0791 | -100.1643
5.0828 | -100.2055
5.0866 | -100.2489
5.0903 | -100.2905
5.094 | -100.3273
5.0978 | -100.3693
5.1015 | -100.4114
5.1053 | -100.442
5.109 | -100.47
5.1128 | -100.4957
5.1165 | -100.5055
5.1202 | -100.5058
5.124 | -100.4921
5.1277 | -100.4684
5.1315 | -100.4361
5.1352 | -100.4078
5.139 | -100.3827
5.1427 | -100.3651
5.1464 | -100.3527
5.1502 | -100.3496
5.1539 | -100.3447
5.1577 | -100.3431
5.1614 | -100.3293
5.1652 | -100.312
5.1689 | -100.2991
5.1726 | -100.2869
5.1764 | -100.2777
5.1801 | -100.2795
5.1839 | -100.2975
5.1876 | -100.3304
5.1914 | -100.3691
5.1951 | -100.4222
5.1988 | -100.4804
5.2026 | -100.5442
5.2063 | -100.6095
5.2101 | -100.6653
5.2138 | -100.711
5.2176 | -100.7494
5.2213 | -100.7754
5.225 | -100.7893
5.2288 | -100.7841
5.2325 | -100.7733
5.2363 | -100.7545
5.24 | -100.744
5.2438 | -100.7329
5.2475 | -100.7196
5.2512 | -100.7161
5.255 | -100.7214
5.2587 | -100.7227
5.2625 | -100.7211
5.2662 | -100.719
5.27 | -100.7141
5.2737 | -100.7023
5.2774 | -100.683
5.2812 | -100.6593
5.2849 | -100.6224
5.2887 | -100.5829
5.2924 | -100.5411
5.2962 | -100.4976
5.2999 | -100.4662
5.3036 | -100.4565
5.3074 | -100.4708
5.3111 | -100.5008
5.3149 | -100.5384
5.3186 | -100.575
5.3224 | -100.6038
5.3261 | -100.6201
5.3298 | -100.6196
5.3336 | -100.6005
5.3373 | -100.5686
5.3411 | -100.5315
5.3448 | -100.493
5.3486 | -100.4448
5.3523 | -100.3892
5.356 | -100.3375
5.3598 | -100.2872
5.3635 | -100.2435
5.3673 | -100.2095
5.371 | -100.1846
5.3748 | -100.1735
5.3785 | -100.1732
5.3822 | -100.1736
5.386 | -100.1887
5.3897 | -100.2201
5.3935 | -100.2633
5.3972 | -100.34
5.401 | -100.3997
5.4047 | -100.4673
5.4084 | -100.5358
5.4122 | -100.6017
5.4159 | -100.6582
5.4197 | -100.7108
5.4234 | -100.7479
5.4272 | -100.7746
5.4309 | -100.809
5.4346 | -100.8491
5.4384 | -100.8912
5.4421 | -100.9393
5.4459 | -100.9943
5.4496 | -101.0473
5.4534 | -101.0925
5.4571 | -101.1309
5.4608 | -101.1477
5.4646 | -101.1425
5.4683 | -101.1042
5.4721 | -101.0512
5.4758 | -100.9813
5.4796 | -100.8898
5.4833 | -100.7904
5.487 | -100.6982
5.4908 | -100.6132
5.4945 | -100.5441
5.4983 | -100.498
5.502 | -100.4834
5.5058 | -100.4902
5.5095 | -100.525
5.5132 | -100.5828
5.517 | -100.6492
5.5207 | -100.7215
5.5245 | -100.7883
5.5282 | -100.8478
5.532 | -100.9092
5.5357 | -100.9618
5.5394 | -101.0081
5.5432 | -101.0405
5.5469 | -101.069
5.5507 | -101.0938
5.5544 | -101.1088
5.5582 | -101.1052
5.5619 | -101.0936
5.5656 | -101.0692
5.5694 | -101.0287
5.5731 | -100.9661
5.5769 | -100.9014
5.5806 | -100.8374
5.5844 | -100.7734
5.5881 | -100.7152
5.5918 | -100.6681
5.5956 | -100.6368
5.5993 | -100.6194
5.6031 | -100.6238
5.6068 | -100.6428
5.6106 | -100.684
5.6143 | -100.7346
5.618 | -100.7925
5.6218 | -100.8693
5.6255 | -100.954
5.6293 | -101.0461
5.633 | -101.1339
5.6368 | -101.2064
5.6405 | -101.2597
5.6442 | -101.2967
5.648 | -101.3167
5.6517 | -101.3227
5.6555 | -101.3181
5.6592 | -101.312
5.663 | -101.2922
5.6667 | -101.2714
5.6704 | -101.2521
5.6742 | -101.2285
5.6779 | -101.2021
5.6817 | -101.1738
5.6854 | -101.1537
5.6892 | -101.1317
5.6929 | -101.1125
5.6966 | -101.1024
5.7004 | -101.1048
5.7041 | -101.1099
5.7079 | -101.1153
5.7116 | -101.127
5.7154 | -101.1343
5.7191 | -101.133
5.7228 | -101.1448
5.7266 | -101.1542
5.7303 | -101.1557
5.7341 | -101.1631
5.7378 | -101.1722
5.7416 | -101.1763
5.7453 | -101.176
5.749 | -101.176
5.7528 | -101.1804
5.7565 | -101.1896
5.7603 | -101.2045
5.764 | -101.2305
5.7678 | -101.2649
5.7715 | -101.3
5.7752 | -101.3451
5.779 | -101.3969
5.7827 | -101.4444
5.7865 | -101.4858
5.7902 | -101.5262
5.794 | -101.566
5.7977 | -101.5974
5.8014 | -101.6238
5.8052 | -101.6528
5.8089 | -101.6662
5.8127 | -101.6792
5.8164 | -101.6807
5.8202 | -101.6792
5.8239 | -101.6778
5.8276 | -101.6754
5.8314 | -101.6741
5.8351 | -101.6713
5.8389 | -101.662
5.8426 | -101.6487
5.8464 | -101.6399
5.8501 | -101.631
5.8538 | -101.6192
5.8576 | -101.618
5.8613 | -101.6171
5.8651 | -101.6198
5.8688 | -101.6209
5.8726 | -101.6259
5.8763 | -101.6243
5.88 | -101.6141
5.8838 | -101.6053
5.8875 | -101.5871
5.8913 | -101.5521
5.895 | -101.5172
5.8988 | -101.4891
5.9025 | -101.469
5.9062 | -101.4543
5.91 | -101.4487
5.9137 | -101.4595
5.9175 | -101.4756
5.9212 | -101.5003
5.925 | -101.5277
5.9287 | -101.5548
5.9324 | -101.5674
5.9362 | -101.5315
5.9399 | -101.5106
5.9437 | -101.4811
5.9474 | -101.4435
5.9512 | -101.4084
5.9549 | -101.3713
5.9586 | -101.3432
5.9624 | -101.3201
5.9661 | -101.3095
5.9699 | -101.3066
5.9736 | -101.3199
5.9774 | -101.3527
5.9811 | -101.4014
5.9848 | -101.458
5.9886 | -101.5193
5.9923 | -101.5855
5.9961 | -101.6477
5.9998 | -101.7083
6.0036 | -101.7515
6.0073 | -101.7838
6.011 | -101.7979
6.0148 | -101.7924
6.0185 | -101.7647
6.0223 | -101.7185
6.026 | -101.6564
6.0298 | -101.584
6.0335 | -101.4971
6.0372 | -101.4043
6.041 | -101.3073
6.0447 | -101.2149
6.0485 | -101.136
6.0522 | -101.0716
6.056 | -101.0252
6.0597 | -100.9975
6.0634 | -100.9888
6.0672 | -100.9933
6.0709 | -101.0153
6.0747 | -101.0586
6.0784 | -101.1153
6.0822 | -101.1766
6.0859 | -101.2383
6.0896 | -101.2902
6.0934 | -101.3332
6.0971 | -101.3714
6.1009 | -101.3914
6.1046 | -101.3994
6.1084 | -101.3935
6.1121 | -101.3804
6.1158 | -101.3625
6.1196 | -101.3345
6.1233 | -101.3033
6.1271 | -101.272
6.1308 | -101.236
6.1346 | -101.202
6.1383 | -101.1693
6.142 | -101.1415
6.1458 | -101.1177
6.1495 | -101.1043
6.1533 | -101.1023
6.157 | -101.1116
6.1608 | -101.1359
6.1645 | -101.1685
6.1682 | -101.2071
6.172 | -101.2528
6.1757 | -101.3121
6.1795 | -101.3788
6.1832 | -101.4468
6.187 | -101.5071
6.1907 | -101.5592
6.1944 | -101.6024
6.1982 | -101.6378
6.2019 | -101.6699
6.2057 | -101.7001
6.2094 | -101.7232
6.2132 | -101.7435
6.2169 | -101.7622
6.2206 | -101.7733
6.2244 | -101.7916
6.2281 | -101.8112
6.2319 | -101.822
6.2356 | -101.8262
6.2394 | -101.8326
6.2431 | -101.8325
6.2468 | -101.8248
6.2506 | -101.8157
6.2543 | -101.8022
6.2581 | -101.7836
6.2618 | -101.7579
6.2656 | -101.7247
6.2693 | -101.6869
6.273 | -101.6475
6.2768 | -101.6164
6.2805 | -101.5904
6.2843 | -101.5742
6.288 | -101.5726
6.2918 | -101.5804
6.2955 | -101.5998
6.2992 | -101.6305
6.303 | -101.6746
6.3067 | -101.7348
6.3105 | -101.7999
6.3142 | -101.8633
6.318 | -101.9184
6.3217 | -101.969
6.3254 | -102.0221
6.3292 | -102.0673
6.3329 | -102.0939
6.3367 | -102.1101
6.3404 | -102.1183
6.3442 | -102.1175
6.3479 | -102.1066
6.3516 | -102.0923
6.3554 | -102.0703
6.3591 | -102.0517
6.3629 | -102.0262
6.3666 | -101.9998
6.3704 | -101.9712
6.3741 | -101.9529
6.3778 | -101.9423
6.3816 | -101.9387
6.3853 | -101.9434
6.3891 | -101.9437
6.3928 | -101.9393
6.3966 | -101.935
6.4003 | -101.9281
6.404 | -101.9169
6.4078 | -101.9112
6.4115 | -101.9082
6.4153 | -101.9025
6.419 | -101.8911
6.4228 | -101.8876
6.4265 | -101.8938
6.4302 | -101.9039
6.434 | -101.9183
6.4377 | -101.9343
6.4415 | -101.9611
6.4452 | -101.9942
6.449 | -102.036
6.4527 | -102.0828
6.4564 | -102.1408
6.4602 | -102.2028
6.4639 | -102.2649
6.4677 | -102.3153
6.4714 | -102.3592
6.4752 | -102.3945
6.4789 | -102.3943
6.4826 | -102.4024
6.4864 | -102.4058
6.4901 | -102.4027
6.4939 | -102.3823
6.4976 | -102.3558
6.5014 | -102.3296
6.5051 | -102.2967
6.5088 | -102.2612
6.5126 | -102.2243
6.5163 | -102.1946
6.5201 | -102.1646
6.5238 | -102.1269
6.5276 | -102.0952
6.5313 | -102.0744
6.535 | -102.0619
6.5388 | -102.0454
6.5425 | -102.0383
6.5463 | -102.0371
6.55 | -102.0352
6.5538 | -102.0339
6.5575 | -102.0391
6.5612 | -102.0544
6.565 | -102.0662
6.5687 | -102.0755
6.5725 | -102.0812
6.5762 | -102.0809
6.58 | -102.0776
6.5837 | -102.0753
6.5874 | -102.0594
6.5912 | -102.0398
6.5949 | -102.0218
6.5987 | -101.9935
6.6024 | -101.9623
6.6062 | -101.9318
6.6099 | -101.9007
6.6136 | -101.8824
6.6174 | -101.8714
6.6211 | -101.8676
6.6249 | -101.8791
6.6286 | -101.8974
6.6324 | -101.9183
6.6361 | -101.9514
6.6398 | -101.9942
6.6436 | -102.04
6.6473 | -102.0813
6.6511 | -102.1135
6.6548 | -102.1422
6.6586 | -102.159
6.6623 | -102.1644
6.666 | -102.1711
6.6698 | -102.1715
6.6735 | -102.1683
6.6773 | -102.1648
6.681 | -102.1617
6.6848 | -102.1595
6.6885 | -102.1649
6.6922 | -102.169
6.696 | -102.1734
6.6997 | -102.1899
6.7035 | -102.2129
6.7072 | -102.2418
6.711 | -102.2842
6.7147 | -102.3327
6.7184 | -102.3861
6.7222 | -102.4475
6.7259 | -102.513
6.7297 | -102.5763
6.7334 | -102.6396
6.7372 | -102.6847
6.7409 | -102.723
6.7446 | -102.7448
6.7484 | -102.7556
6.7521 | -102.7584
6.7559 | -102.7445
6.7596 | -102.722
6.7634 | -102.6855
6.7671 | -102.6429
6.7708 | -102.604
6.7746 | -102.5647
6.7783 | -102.5246
6.7821 | -102.4765
6.7858 | -102.4308
6.7896 | -102.3899
6.7933 | -102.3532
6.797 | -102.3217
6.8008 | -102.2859
6.8045 | -102.2565
6.8083 | -102.2342
6.812 | -102.2121
6.8158 | -102.1983
6.8195 | -102.1969
6.8232 | -102.2024
6.827 | -102.2094
6.8307 | -102.2222
6.8345 | -102.2339
6.8382 | -102.2584
6.842 | -102.2805
6.8457 | -102.3023
6.8494 | -102.3268
6.8532 | -102.3573
6.8569 | -102.3845
6.8607 | -102.4083
6.8644 | -102.4354
6.8682 | -102.4445
6.8719 | -102.4401
6.8756 | -102.4377
6.8794 | -102.4273
6.8831 | -102.4157
6.8869 | -102.3966
6.8906 | -102.3732
6.8944 | -102.3439
6.8981 | -102.3166
6.9018 | -102.3002
6.9056 | -102.2864
6.9093 | -102.2787
6.9131 | -102.285
6.9168 | -102.2885
6.9206 | -102.2988
6.9243 | -102.3184
6.928 | -102.3428
6.9318 | -102.366
6.9355 | -102.3896
6.9393 | -102.3997
6.943 | -102.401
6.9468 | -102.3887
6.9505 | -102.3693
6.9542 | -102.3444
6.958 | -102.3269
6.9617 | -102.3134
6.9655 | -102.2999
6.9692 | -102.296
6.973 | -102.2952
6.9767 | -102.3021
6.9804 | -102.3184
6.9842 | -102.3351
6.9879 | -102.3565
6.9917 | -102.3846
6.9954 | -102.4207
6.9992 | -102.4559
7.0029 | -102.4824
7.0066 | -102.5089
7.0104 | -102.5355
7.0141 | -102.555
7.0179 | -102.5699
7.0216 | -102.5986
7.0254 | -102.6215
7.0291 | -102.637
7.0328 | -102.6516
7.0366 | -102.6671
7.0403 | -102.6855
7.0441 | -102.7112
7.0478 | -102.7395
7.0516 | -102.7744
7.0553 | -102.8021
7.059 | -102.8224
7.0628 | -102.843
7.0665 | -102.8664
7.0703 | -102.8915
7.074 | -102.9167
7.0778 | -102.9398
7.0815 | -102.9537
7.0852 | -102.9678
7.089 | -102.9876
7.0927 | -103.0012
7.0965 | -103.0058
7.1002 | -103.0047
7.104 | -103.0003
7.1077 | -102.9924
7.1114 | -102.9738
7.1152 | -102.9487
7.1189 | -102.9216
7.1227 | -102.897
7.1264 | -102.8685
7.1302 | -102.8371
7.1339 | -102.8109
7.1376 | -102.7863
7.1414 | -102.7619
7.1451 | -102.7351
7.1489 | -102.7119
7.1526 | -102.6941
7.1564 | -102.6808
7.1601 | -102.6721
7.1638 | -102.6679
7.1676 | -102.6686
7.1713 | -102.6657
7.1751 | -102.6668
7.1788 | -102.67
7.1826 | -102.6704
7.1863 | -102.6681
7.19 | -102.6657
7.1938 | -102.6579
7.1975 | -102.6431
7.2013 | -102.6259
7.205 | -102.6036
7.2088 | -102.5816
7.2125 | -102.5596
7.2162 | -102.5358
7.22 | -102.5213
7.2237 | -102.5075
7.2275 | -102.4923
7.2312 | -102.4785
7.235 | -102.4669
7.2387 | -102.4598
7.2424 | -102.4587
7.2462 | -102.4621
7.2499 | -102.4706
7.2537 | -102.4853
7.2574 | -102.5018
7.2612 | -102.5173
7.2649 | -102.5347
7.2686 | -102.5562
7.2724 | -102.5829
7.2761 | -102.6043
7.2799 | -102.6275
7.2836 | -102.6457
7.2874 | -102.656
7.2911 | -102.6572
7.2948 | -102.6514
7.2986 | -102.6334
7.3023 | -102.6087
7.3061 | -102.5754
7.3098 | -102.5359
7.3136 | -102.4969
7.3173 | -102.4581
7.321 | -102.4265
7.3248 | -102.4088
7.3285 | -102.3973
7.3323 | -102.4034
7.336 | -102.4284
7.3398 | -102.4674
7.3435 | -102.528
7.3472 | -102.6018
7.351 | -102.6824
7.3547 | -102.7696
7.3585 | -102.859
7.3622 | -102.9445
7.366 | -103.0244
7.3697 | -103.0968
7.3734 | -103.1582
7.3772 | -103.2034
7.3809 | -103.2302
7.3847 | -103.2485
7.3884 | -103.248
7.3922 | -103.238
7.3959 | -103.2223
7.3996 | -103.2088
7.4034 | -103.1907
7.4071 | -103.1761
7.4109 | -103.1591
7.4146 | -103.1383
7.4184 | -103.1162
7.4221 | -103.0905
7.4258 | -103.0586
7.4296 | -103.0235
7.4333 | -102.9837
7.4371 | -102.945
7.4408 | -102.9036
7.4446 | -102.859
7.4483 | -102.816
7.452 | -102.7746
7.4558 | -102.7398
7.4595 | -102.7131
7.4633 | -102.6951
7.467 | -102.689
7.4708 | -102.6928
7.4745 | -102.7005
7.4782 | -102.7171
7.482 | -102.7412
7.4857 | -102.7652
7.4895 | -102.7931
7.4932 | -102.8179
7.497 | -102.8411
7.5007 | -102.8609
7.5044 | -102.8777
7.5082 | -102.8876
7.5119 | -102.8882
7.5157 | -102.8842
7.5194 | -102.8812
7.5232 | -102.879
7.5269 | -102.8752
7.5306 | -102.8679
7.5344 | -102.8663
7.5381 | -102.8709
7.5419 | -102.8778
7.5456 | -102.8879
7.5494 | -102.9008
7.5531 | -102.918
7.5568 | -102.935
7.5606 | -102.9535
7.5643 | -102.9684
7.5681 | -102.981
7.5718 | -102.987
7.5756 | -102.9957
7.5793 | -103.0008
7.583 | -103.0002
7.5868 | -102.9997
7.5905 | -102.9957
7.5943 | -102.9876
7.598 | -102.9798
7.6018 | -102.9696
7.6055 | -102.9576
7.6092 | -102.9444
7.613 | -102.9327
7.6167 | -102.9217
7.6205 | -102.9063
7.6242 | -102.8975
7.628 | -102.8851
7.6317 | -102.8752
7.6354 | -102.8634
7.6392 | -102.8527
7.6429 | -102.8467
7.6467 | -102.8336
7.6504 | -102.819
7.6542 | -102.8152
7.6579 | -102.815
7.6616 | -102.8163
7.6654 | -102.8235
7.6691 | -102.8331
7.6729 | -102.8453
7.6766 | -102.8583
7.6804 | -102.8786
7.6841 | -102.9075
7.6878 | -102.9407
7.6916 | -102.9689
7.6953 | -103.0018
7.6991 | -103.0407
7.7028 | -103.0804
7.7066 | -103.1182
7.7103 | -103.1549
7.714 | -103.1973
7.7178 | -103.2389
7.7215 | -103.2756
7.7253 | -103.3089
7.729 | -103.3392
7.7328 | -103.3674
7.7365 | -103.3941
7.7402 | -103.4129
7.744 | -103.4335
7.7477 | -103.4433
7.7515 | -103.4477
7.7552 | -103.4469
7.759 | -103.4385
7.7627 | -103.4346
7.7664 | -103.4301
7.7702 | -103.414
7.7739 | -103.3936
7.7777 | -103.3762
7.7814 | -103.3601
7.7852 | -103.3466
7.7889 | -103.3315
7.7926 | -103.3243
7.7964 | -103.3151
7.8001 | -103.3005
7.8039 | -103.2907
7.8076 | -103.2891
7.8114 | -103.2914
7.8151 | -103.3002
7.8188 | -103.3122
7.8226 | -103.3311
7.8263 | -103.3559
7.8301 | -103.3788
7.8338 | -103.4043
7.8376 | -103.4361
7.8413 | -103.4683
7.845 | -103.4965
7.8488 | -103.526
7.8525 | -103.5484
7.8563 | -103.5683
7.86 | -103.5815
7.8638 | -103.59
7.8675 | -103.5942
7.8712 | -103.5918
7.875 | -103.5842
7.8787 | -103.5684
7.8825 | -103.545
7.8862 | -103.5212
7.89 | -103.4996
7.8937 | -103.4764
7.8974 | -103.4494
7.9012 | -103.4211
7.9049 | -103.3967
7.9087 | -103.3768
7.9124 | -103.3643
7.9162 | -103.3577
7.9199 | -103.3531
7.9236 | -103.3494
7.9274 | -103.3481
7.9311 | -103.348
7.9349 | -103.3533
7.9386 | -103.3612
7.9424 | -103.3734
7.9461 | -103.3855
7.9498 | -103.4023
7.9536 | -103.421
7.9573 | -103.4379
7.9611 | -103.4516
7.9648 | -103.46
7.9686 | -103.475
7.9723 | -103.489
7.976 | -103.5002
7.9798 | -103.5148
7.9835 | -103.5256
7.9873 | -103.5337
7.991 | -103.5399
7.9948 | -103.5474
7.9985 | -103.5529
8.0022 | -103.559
8.006 | -103.5628
8.0097 | -103.5698
8.0135 | -103.5773
8.0172 | -103.582
8.021 | -103.5853
8.0247 | -103.5833
8.0284 | -103.5785
8.0322 | -103.5805
8.0359 | -103.583
8.0397 | -103.5801
8.0434 | -103.5731
8.0472 | -103.562
8.0509 | -103.5501
8.0546 | -103.5315
8.0584 | -103.5085
8.0621 | -103.4914
8.0659 | -103.4744
8.0696 | -103.4563
8.0734 | -103.4414
8.0771 | -103.4323
8.0808 | -103.4271
8.0846 | -103.4136
8.0883 | -103.3981
8.0921 | -103.3845
8.0958 | -103.3685
8.0996 | -103.3569
8.1033 | -103.3371
8.107 | -103.3197
8.1108 | -103.3055
8.1145 | -103.2891
8.1183 | -103.2759
8.122 | -103.265
8.1258 | -103.2573
8.1295 | -103.2506
8.1332 | -103.2467
8.137 | -103.2494
8.1407 | -103.2504
8.1445 | -103.2475
8.1482 | -103.2441
8.152 | -103.2362
8.1557 | -103.2261
8.1594 | -103.2197
8.1632 | -103.2148
8.1669 | -103.2164
8.1707 | -103.2167
8.1744 | -103.2201
8.1782 | -103.2284
8.1819 | -103.2383
8.1856 | -103.2483
8.1894 | -103.26
8.1931 | -103.2734
8.1969 | -103.2909
8.2006 | -103.3088
8.2044 | -103.3308
8.2081 | -103.3521
8.2118 | -103.3699
8.2156 | -103.3833
8.2193 | -103.3978
8.2231 | -103.411
8.2268 | -103.4228
8.2306 | -103.4207
8.2343 | -103.4144
8.238 | -103.4061
8.2418 | -103.3947
8.2455 | -103.3818
8.2493 | -103.3687
8.253 | -103.3583
8.2568 | -103.3488
8.2605 | -103.3386
8.2642 | -103.3332
8.268 | -103.3368
8.2717 | -103.3434
8.2755 | -103.3546
8.2792 | -103.3735
8.283 | -103.3926
8.2867 | -103.4152
8.2904 | -103.4388
8.2942 | -103.4666
8.2979 | -103.4913
8.3017 | -103.511
8.3054 | -103.5298
8.3092 | -103.5516
8.3129 | -103.5774
8.3166 | -103.6014
8.3204 | -103.6215
8.3241 | -103.6396
8.3279 | -103.6567
8.3316 | -103.674
8.3354 | -103.6919
8.3391 | -103.7098
8.3428 | -103.7247
8.3466 | -103.7371
8.3503 | -103.752
8.3541 | -103.769
8.3578 | -103.7785
8.3616 | -103.7796
8.3653 | -103.778
8.369 | -103.772
8.3728 | -103.7674
8.3765 | -103.7595
8.3803 | -103.7489
8.384 | -103.7381
8.3878 | -103.7268
8.3915 | -103.7124
8.3952 | -103.6999
8.399 | -103.6849
8.4027 | -103.6694
8.4065 | -103.6588
8.4102 | -103.6559
8.414 | -103.6562
8.4177 | -103.6572
8.4214 | -103.6664
8.4252 | -103.6809
8.4289 | -103.6961
8.4327 | -103.7103
8.4364 | -103.7155
8.4402 | -103.7238
8.4439 | -103.733
8.4476 | -103.7413
8.4514 | -103.7493
8.4551 | -103.75
8.4589 | -103.7478
8.4626 | -103.7405
8.4664 | -103.7295
8.4701 | -103.7174
8.4738 | -103.7042
8.4776 | -103.6883
8.4813 | -103.6716
8.4851 | -103.6633
8.4888 | -103.657
8.4926 | -103.653
8.4963 | -103.6533
8.5 | -103.6593
8.5038 | -103.6734
8.5075 | -103.6912
8.5113 | -103.7099
8.515 | -103.7327
8.5188 | -103.7579
8.5225 | -103.7835
8.5262 | -103.8056
8.53 | -103.8227
8.5337 | -103.8287
8.5375 | -103.8335
8.5412 | -103.8367
8.545 | -103.8349
8.5487 | -103.8308
8.5524 | -103.8242
8.5562 | -103.8158
8.5599 | -103.811
8.5637 | -103.804
8.5674 | -103.7967
8.5712 | -103.7919
8.5749 | -103.7878
8.5786 | -103.7927
8.5824 | -103.802
8.5861 | -103.8148
8.5899 | -103.8282
8.5936 | -103.8426
8.5974 | -103.8604
8.6011 | -103.8827
8.6048 | -103.9069
8.6086 | -103.9249
8.6123 | -103.9383
8.6161 | -103.9519
8.6198 | -103.9596
8.6236 | -103.9638
8.6273 | -103.967
8.631 | -103.9646
8.6348 | -103.959
8.6385 | -103.9474
8.6423 | -103.9427
8.646 | -103.9386
8.6498 | -103.9338
8.6535 | -103.9327
8.6572 | -103.9361
8.661 | -103.9352
8.6647 | -103.9357
8.6685 | -103.9399
8.6722 | -103.9416
8.676 | -103.9416
8.6797 | -103.9412
8.6834 | -103.9454
8.6872 | -103.9551
8.6909 | -103.9635
8.6947 | -103.9727
8.6984 | -103.9798
8.7022 | -103.9869
8.7059 | -103.9957
8.7096 | -104.0064
8.7134 | -104.0188
8.7171 | -104.0336
8.7209 | -104.0524
8.7246 | -104.0713
8.7284 | -104.09
8.7321 | -104.1082
8.7358 | -104.1297
8.7396 | -104.1542
8.7433 | -104.1737
8.7471 | -104.1957
8.7508 | -104.2207
8.7546 | -104.2428
8.7583 | -104.2654
8.762 | -104.2862
8.7658 | -104.3067
8.7695 | -104.3197
8.7733 | -104.326
8.777 | -104.3296
8.7808 | -104.3333
8.7845 | -104.332
8.7882 | -104.3223
8.792 | -104.3194
8.7957 | -104.308
8.7995 | -104.2937
8.8032 | -104.2857
8.807 | -104.2754
8.8107 | -104.2652
8.8144 | -104.2488
8.8182 | -104.2281
8.8219 | -104.2112
8.8257 | -104.1919
8.8294 | -104.1707
8.8332 | -104.1453
8.8369 | -104.1156
8.8406 | -104.0856
8.8444 | -104.0594
8.8481 | -104.0333
8.8519 | -104.008
8.8556 | -103.981
8.8594 | -103.9591
8.8631 | -103.9416
8.8668 | -103.9281
8.8706 | -103.9182
8.8743 | -103.9157
8.8781 | -103.9116
8.8818 | -103.9082
8.8856 | -103.9084
8.8893 | -103.9131
8.893 | -103.9222
8.8968 | -103.9314
8.9005 | -103.9421
8.9043 | -103.9514
8.908 | -103.9563
8.9118 | -103.9549
8.9155 | -103.9558
8.9192 | -103.9554
8.923 | -103.9465
8.9267 | -103.9381
8.9305 | -103.9321
8.9342 | -103.9221
8.938 | -103.9113
8.9417 | -103.9025
8.9454 | -103.8972
8.9492 | -103.8939
8.9529 | -103.8905
8.9567 | -103.8942
8.9604 | -103.9003
8.9642 | -103.9078
8.9679 | -103.9172
8.9716 | -103.9224
8.9754 | -103.9243
8.9791 | -103.9338
8.9829 | -103.942
8.9866 | -103.9536
8.9904 | -103.9652
8.9941 | -103.9778
8.9978 | -103.9842
9.0016 | -103.9905
9.0053 | -103.9966
9.0091 | -104.0026
9.0128 | -104.0031
9.0166 | -104.0031
9.0203 | -103.9985
9.024 | -103.9887
9.0278 | -103.9798
9.0315 | -103.9724
9.0353 | -103.9661
9.039 | -103.9562
9.0428 | -103.9486
9.0465 | -103.9385
9.0502 | -103.9282
9.054 | -103.9224
9.0577 | -103.9154
9.0615 | -103.9061
9.0652 | -103.9028
9.069 | -103.9043
9.0727 | -103.9044
9.0764 | -103.9046
9.0802 | -103.9055
9.0839 | -103.9098
9.0877 | -103.9133
9.0914 | -103.9191
9.0952 | -103.9252
9.0989 | -103.9327
9.1026 | -103.946
9.1064 | -103.9556
9.1101 | -103.9588
9.1139 | -103.964
9.1176 | -103.9711
9.1214 | -103.9776
9.1251 | -103.9833
9.1288 | -103.9938
9.1326 | -104.0044
9.1363 | -104.0144
9.1401 | -104.0239
9.1438 | -104.0396
9.1476 | -104.0577
9.1513 | -104.0793
9.155 | -104.1047
9.1588 | -104.1305
9.1625 | -104.1538
9.1663 | -104.1794
9.17 | -104.2093
9.1738 | -104.2364
9.1775 | -104.2521
9.1812 | -104.2721
9.185 | -104.2881
9.1887 | -104.3021
9.1925 | -104.3102
9.1962 | -104.3152
9.2 | -104.3155
9.2037 | -104.3145
9.2074 | -104.3126
9.2112 | -104.3098
9.2149 | -104.3053
9.2187 | -104.2893
9.2224 | -104.2724
9.2262 | -104.2553
9.2299 | -104.2361
9.2336 | -104.2145
9.2374 | -104.1828
9.2411 | -104.1551
9.2449 | -104.1246
9.2486 | -104.0901
9.2524 | -104.0494
9.2561 | -104.0067
9.2598 | -103.9668
9.2636 | -103.9263
9.2673 | -103.8851
9.2711 | -103.85
9.2748 | -103.8202
9.2786 | -103.793
9.2823 | -103.7721
9.286 | -103.7571
9.2898 | -103.752
9.2935 | -103.7515
9.2973 | -103.7611
9.301 | -103.7769
9.3048 | -103.8008
9.3085 | -103.8304
9.3122 | -103.8692
9.316 | -103.9099
9.3197 | -103.9514
9.3235 | -103.9977
9.3272 | -104.0401
9.331 | -104.0839
9.3347 | -104.1203
9.3384 | -104.1517
9.3422 | -104.1767
9.3459 | -104.1982
9.3497 | -104.2128
9.3534 | -104.2209
9.3572 | -104.2321
9.3609 | -104.2353
9.3646 | -104.2371
9.3684 | -104.24
9.3721 | -104.2412
9.3759 | -104.2524
9.3796 | -104.259
9.3834 | -104.2632
9.3871 | -104.2727
9.3908 | -104.2863
9.3946 | -104.3024
9.3983 | -104.3176
9.4021 | -104.3324
9.4058 | -104.3509
9.4096 | -104.3639
9.4133 | -104.3774
9.417 | -104.3921
9.4208 | -104.4027
9.4245 | -104.4129
9.4283 | -104.4184
9.432 | -104.4204
9.4358 | -104.4216
9.4395 | -104.4192
9.4432 | -104.4164
9.447 | -104.4135
9.4507 | -104.4018
9.4545 | -104.3891
9.4582 | -104.3785
9.462 | -104.3654
9.4657 | -104.3526
9.4694 | -104.3387
9.4732 | -104.3283
9.4769 | -104.3185
9.4807 | -104.3105
9.4844 | -104.3068
9.4882 | -104.308
9.4919 | -104.3111
9.4956 | -104.3137
9.4994 | -104.3236
9.5031 | -104.3386
9.5069 | -104.353
9.5106 | -104.3727
9.5144 | -104.3908
9.5181 | -104.4176
9.5218 | -104.4496
9.5256 | -104.4833
9.5293 | -104.5179
9.5331 | -104.5552
9.5368 | -104.5912
9.5406 | -104.6286
9.5443 | -104.6599
9.548 | -104.6948
9.5518 | -104.7254
9.5555 | -104.7478
9.5593 | -104.766
9.563 | -104.7811
9.5668 | -104.7959
9.5705 | -104.8088
9.5742 | -104.8123
9.578 | -104.8145
9.5817 | -104.81
9.5855 | -104.8017
9.5892 | -104.7931
9.593 | -104.7788
9.5967 | -104.7559
9.6004 | -104.7296
9.6042 | -104.7028
9.6079 | -104.6729
9.6117 | -104.6368
9.6154 | -104.5999
9.6192 | -104.5633
9.6229 | -104.52
9.6266 | -104.4763
9.6304 | -104.44
9.6341 | -104.4092
9.6379 | -104.3767
9.6416 | -104.3469
9.6454 | -104.3219
9.6491 | -104.3011
9.6528 | -104.29
9.6566 | -104.2768
9.6603 | -104.2702
9.6641 | -104.2656
9.6678 | -104.2671
9.6716 | -104.271
9.6753 | -104.2825
9.679 | -104.298
9.6828 | -104.3169
9.6865 | -104.3338
9.6903 | -104.3508
9.694 | -104.3686
9.6978 | -104.3869
9.7015 | -104.4022
9.7052 | -104.4124
9.709 | -104.4213
9.7127 | -104.425
9.7165 | -104.4236
9.7202 | -104.4207
9.724 | -104.4203
9.7277 | -104.4206
9.7314 | -104.4216
9.7352 | -104.4246
9.7389 | -104.4278
9.7427 | -104.4303
9.7464 | -104.4286
9.7502 | -104.4263
9.7539 | -104.4258
9.7576 | -104.4266
9.7614 | -104.4251
9.7651 | -104.4258
9.7689 | -104.4275
9.7726 | -104.424
9.7764 | -104.4205
9.7801 | -104.4183
9.7838 | -104.4135
9.7876 | -104.4092
9.7913 | -104.4037
9.7951 | -104.3975
9.7988 | -104.3919
9.8026 | -104.382
9.8063 | -104.3734
9.8101 | -104.3639
9.8138 | -104.3558
9.8175 | -104.3479
9.8213 | -104.3388
9.825 | -104.3308
9.8288 | -104.3277
9.8325 | -104.3294
9.8362 | -104.3338
9.84 | -104.3367
9.8437 | -104.3458
9.8475 | -104.3605
9.8512 | -104.3805
9.855 | -104.4032
9.8587 | -104.4291
9.8625 | -104.4592
9.8662 | -104.496
9.8699 | -104.5272
9.8737 | -104.5555
9.8774 | -104.5845
9.8812 | -104.6138
9.8849 | -104.6451
9.8887 | -104.6718
9.8924 | -104.6973
9.8961 | -104.7232
9.8999 | -104.7465
9.9036 | -104.7608
9.9074 | -104.7754
9.9111 | -104.7855
9.9148 | -104.7884
9.9186 | -104.7896
9.9223 | -104.7874
9.9261 | -104.7819
9.9298 | -104.7739
9.9336 | -104.7646
9.9373 | -104.7534
9.9411 | -104.7379
9.9448 | -104.7241
9.9485 | -104.7124
9.9523 | -104.7005
9.956 | -104.6894
9.9598 | -104.6814
9.9635 | -104.6775
9.9672 | -104.6774
9.971 | -104.6757
9.9747 | -104.6747
9.9785 | -104.6792
9.9822 | -104.6822
9.986 | -104.6915
9.9897 | -104.7014
9.9934 | -104.7091
9.9972 | -104.7179
10.0009 | -104.7279
10.0047 | -104.7353
10.0084 | -104.7403
10.0122 | -104.7435
10.0159 | -104.7475
10.0197 | -104.7534
10.0234 | -104.7529
10.0271 | -104.7451
10.0309 | -104.7366
10.0346 | -104.7284
10.0384 | -104.724
10.0421 | -104.7192
10.0458 | -104.7147
10.0496 | -104.7126
10.0533 | -104.7041
10.0571 | -104.7018
10.0608 | -104.7042
10.0646 | -104.7046
10.0683 | -104.7048
10.0721 | -104.7081
10.0758 | -104.7113
10.0795 | -104.7142
10.0833 | -104.7151
10.087 | -104.7165
10.0908 | -104.7204
10.0945 | -104.7246
10.0983 | -104.729
10.102 | -104.7326
10.1057 | -104.735
10.1095 | -104.7422
10.1132 | -104.7509
10.117 | -104.763
10.1207 | -104.7756
10.1244 | -104.7882
10.1282 | -104.8078
10.1319 | -104.8289
10.1357 | -104.8494
10.1394 | -104.8717
10.1432 | -104.8915
10.1469 | -104.9104
10.1507 | -104.9269
10.1544 | -104.9398
10.1581 | -104.9538
10.1619 | -104.9605
10.1656 | -104.9614
10.1694 | -104.9624
10.1731 | -104.9605
10.1768 | -104.9508
10.1806 | -104.937
10.1843 | -104.9169
10.1881 | -104.8941
10.1918 | -104.8727
10.1956 | -104.8526
10.1993 | -104.8328
10.203 | -104.8059
10.2068 | -104.7773
10.2105 | -104.7544
10.2143 | -104.7299
10.218 | -104.7015
10.2218 | -104.6783
10.2255 | -104.6593
10.2293 | -104.6424
10.233 | -104.6264
10.2367 | -104.61
10.2405 | -104.5909
10.2442 | -104.5749
10.248 | -104.5621
10.2517 | -104.5519
10.2555 | -104.5392
10.2592 | -104.5259
10.2629 | -104.5153
10.2667 | -104.5083
10.2704 | -104.5017
10.2742 | -104.4935
10.2779 | -104.4838
10.2817 | -104.473
10.2854 | -104.4648
10.2891 | -104.4591
10.2929 | -104.4569
10.2966 | -104.4592
10.3004 | -104.462
10.3041 | -104.468
10.3079 | -104.4738
10.3116 | -104.4869
10.3153 | -104.5008
10.3191 | -104.5142
10.3228 | -104.5318
10.3266 | -104.5558
10.3303 | -104.5818
10.3341 | -104.6061
10.3378 | -104.6323
10.3415 | -104.6602
10.3453 | -104.6883
10.349 | -104.7123
10.3528 | -104.7333
10.3565 | -104.7514
10.3603 | -104.7643
10.364 | -104.7759
10.3677 | -104.7866
10.3715 | -104.7902
10.3752 | -104.7904
10.379 | -104.7893
10.3827 | -104.7885
10.3865 | -104.7854
10.3902 | -104.7849
10.3939 | -104.7846
10.3977 | -104.7855
10.4014 | -104.7849
10.4052 | -104.7851
10.4089 | -104.7898
10.4127 | -104.7969
10.4164 | -104.8026
10.4201 | -104.807
10.4239 | -104.813
10.4276 | -104.8218
10.4314 | -104.8341
10.4351 | -104.8488
10.4389 | -104.8636
10.4426 | -104.8752
10.4463 | -104.8881
10.4501 | -104.9014
10.4538 | -104.9131
10.4576 | -104.9256
10.4613 | -104.9403
10.4651 | -104.9555
10.4688 | -104.9719
10.4725 | -104.9891
10.4763 | -105.0039
10.48 | -105.0165
10.4838 | -105.0289
10.4875 | -105.0423
10.4913 | -105.0589
10.495 | -105.0757
10.4987 | -105.0932
10.5025 | -105.1163
10.5062 | -105.1402
10.51 | -105.1596
10.5137 | -105.1771
10.5175 | -105.1954
10.5212 | -105.2142
10.5249 | -105.232
10.5287 | -105.2463
10.5324 | -105.2572
10.5362 | -105.2621
10.5399 | -105.2688
10.5437 | -105.2734
10.5474 | -105.2773
10.5511 | -105.2774
10.5549 | -105.2785
10.5586 | -105.2787
10.5624 | -105.2736
10.5661 | -105.2701
10.5699 | -105.2697
10.5736 | -105.2685
10.5773 | -105.2646
10.5811 | -105.2628
10.5848 | -105.2615
10.5886 | -105.2537
10.5923 | -105.2425
10.5961 | -105.2341
10.5998 | -105.2242
10.6035 | -105.2085
10.6073 | -105.1969
10.611 | -105.1839
10.6148 | -105.1676
10.6185 | -105.1519
10.6223 | -105.136
10.626 | -105.1197
10.6297 | -105.1041
10.6335 | -105.0895
10.6372 | -105.0805
10.641 | -105.0705
10.6447 | -105.0609
10.6485 | -105.055
10.6522 | -105.052
10.6559 | -105.0493
10.6597 | -105.0532
10.6634 | -105.0659
10.6672 | -105.0791
10.6709 | -105.0915
10.6747 | -105.1
10.6784 | -105.1075
10.6821 | -105.1183
10.6859 | -105.1306
10.6896 | -105.1404
10.6934 | -105.1419
10.6971 | -105.1391
10.7009 | -105.1313
10.7046 | -105.1193
10.7083 | -105.1025
10.7121 | -105.0838
10.7158 | -105.0636
10.7196 | -105.0348
10.7233 | -104.9995
10.7271 | -104.9659
10.7308 | -104.9375
10.7345 | -104.9082
10.7383 | -104.8782
10.742 | -104.8532
10.7458 | -104.832
10.7495 | -104.8122
10.7533 | -104.7903
10.757 | -104.7735
10.7607 | -104.7577
10.7645 | -104.7438
10.7682 | -104.7334
10.772 | -104.7222
10.7757 | -104.7149
10.7795 | -104.7121
10.7832 | -104.7087
10.7869 | -104.709
10.7907 | -104.7117
10.7944 | -104.7179
10.7982 | -104.7242
10.8019 | -104.7271
10.8057 | -104.7322
10.8094 | -104.742
10.8131 | -104.7557
10.8169 | -104.7627
10.8206 | -104.7696
10.8244 | -104.7826
10.8281 | -104.801
10.8319 | -104.8222
10.8356 | -104.8435
10.8393 | -104.8619
10.8431 | -104.8782
10.8468 | -104.8952
10.8506 | -104.918
10.8543 | -104.9422
10.8581 | -104.9626
10.8618 | -104.9853
10.8655 | -105.0081
10.8693 | -105.0291
10.873 | -105.0471
10.8768 | -105.0647
10.8805 | -105.0836
10.8843 | -105.1012
10.888 | -105.1102
10.8917 | -105.1199
10.8955 | -105.1279
10.8992 | -105.1343
10.903 | -105.1381
10.9067 | -105.1382
10.9105 | -105.1367
10.9142 | -105.1295
10.9179 | -105.1157
10.9217 | -105.1033
10.9254 | -105.0911
10.9292 | -105.0771
10.9329 | -105.0614
10.9367 | -105.0455
10.9404 | -105.0331
10.9441 | -105.015
10.9479 | -105
10.9516 | -104.9923
10.9554 | -104.9894
10.9591 | -104.9864
10.9629 | -104.9836
10.9666 | -104.981
10.9703 | -104.9818
10.9741 | -104.986
10.9778 | -104.9943
10.9816 | -104.9995
10.9853 | -105.0038
10.9891 | -105.0064
10.9928 | -105.0084
10.9965 | -105.0124
11.0003 | -105.0152
11.004 | -105.0188
11.0078 | -105.0219
11.0115 | -105.0197
11.0153 | -105.0195
11.019 | -105.021
11.0227 | -105.0189
11.0265 | -105.0179
11.0302 | -105.0158
11.034 | -105.0152
11.0377 | -105.0144
11.0415 | -105.009
11.0452 | -105.0004
11.0489 | -104.9948
11.0527 | -104.9887
11.0564 | -104.9871
11.0602 | -104.988
11.0639 | -104.9897
11.0677 | -104.9942
11.0714 | -105.0009
11.0751 | -105.0106
11.0789 | -105.0237
11.0826 | -105.0335
11.0864 | -105.0441
11.0901 | -105.0593
11.0939 | -105.0745
11.0976 | -105.0904
11.1013 | -105.1072
11.1051 | -105.1289
11.1088 | -105.1468
11.1126 | -105.1622
11.1163 | -105.1803
11.1201 | -105.2032
11.1238 | -105.2253
11.1275 | -105.2467
11.1313 | -105.2649
11.135 | -105.2805
11.1388 | -105.2993
11.1425 | -105.3219
11.1463 | -105.3409
11.15 | -105.3579
11.1537 | -105.3764
11.1575 | -105.3931
11.1612 | -105.4076
11.165 | -105.4214
11.1687 | -105.4336
11.1725 | -105.4444
11.1762 | -105.4535
11.1799 | -105.4611
11.1837 | -105.4684
11.1874 | -105.4756
11.1912 | -105.4838
11.1949 | -105.4913
11.1987 | -105.4956
11.2024 | -105.4979
11.2061 | -105.4979
11.2099 | -105.4941
11.2136 | -105.4902
11.2174 | -105.4818
11.2211 | -105.4698
11.2249 | -105.4591
11.2286 | -105.4433
11.2323 | -105.4276
11.2361 | -105.4123
11.2398 | -105.3967
11.2436 | -105.3825
11.2473 | -105.366
11.2511 | -105.3485
11.2548 | -105.3342
11.2585 | -105.3211
11.2623 | -105.312
11.266 | -105.3054
11.2698 | -105.3041
11.2735 | -105.3063
11.2773 | -105.3095
11.281 | -105.3128
11.2847 | -105.3205
11.2885 | -105.3314
11.2922 | -105.3443
11.296 | -105.3618
11.2997 | -105.3729
11.3035 | -105.3825
11.3072 | -105.3961
11.3109 | -105.4106
11.3147 | -105.4236
11.3184 | -105.4282
11.3222 | -105.4312
11.3259 | -105.4306
11.3297 | -105.4275
11.3334 | -105.4204
11.3371 | -105.3885
11.3409 | -105.3701
11.3446 | -105.3483
11.3484 | -105.3235
11.3521 | -105.2951
11.3559 | -105.2638
11.3596 | -105.2307
11.3633 | -105.2013
11.3671 | -105.1755
11.3708 | -105.1473
11.3746 | -105.1207
11.3783 | -105.0963
11.3821 | -105.0758
11.3858 | -105.0562
11.3895 | -105.0373
11.3933 | -105.0258
11.397 | -105.0187
11.4008 | -105.013
11.4045 | -105.0058
11.4083 | -105.0008
11.412 | -105.0039
11.4157 | -105.0111
11.4195 | -105.0178
11.4232 | -105.0253
11.427 | -105.0333
11.4307 | -105.0426
11.4345 | -105.0548
11.4382 | -105.0686
11.4419 | -105.0852
11.4457 | -105.1033
11.4494 | -105.1188
11.4532 | -105.1316
11.4569 | -105.1444
11.4607 | -105.1573
11.4644 | -105.1732
11.4681 | -105.1868
11.4719 | -105.2001
11.4756 | -105.214
11.4794 | -105.2292
11.4831 | -105.2492
11.4869 | -105.2665
11.4906 | -105.2799
11.4943 | -105.2927
11.4981 | -105.3042
11.5018 | -105.3179
11.5056 | -105.3342
11.5093 | -105.3445
11.5131 | -105.3553
11.5168 | -105.3639
11.5205 | -105.3698
11.5243 | -105.3733
11.528 | -105.3757
11.5318 | -105.3744
11.5355 | -105.3667
11.5393 | -105.3571
11.543 | -105.3455
11.5467 | -105.3294
11.5505 | -105.3129
11.5542 | -105.2963
11.558 | -105.2785
11.5617 | -105.2579
11.5655 | -105.2335
11.5692 | -105.2098
11.5729 | -105.1884
11.5767 | -105.1676
11.5804 | -105.1448
11.5842 | -105.1271
11.5879 | -105.1126
11.5917 | -105.1004
11.5954 | -105.0915
11.5991 | -105.0889
11.6029 | -105.0879
11.6066 | -105.0907
11.6104 | -105.0955
11.6141 | -105.1046
11.6179 | -105.1184
11.6216 | -105.1342
11.6253 | -105.1528
11.6291 | -105.1701
11.6328 | -105.1903
11.6366 | -105.211
11.6403 | -105.228
11.6441 | -105.2448
11.6478 | -105.2587
11.6515 | -105.2738
11.6553 | -105.2905
11.659 | -105.3061
11.6628 | -105.3194
11.6665 | -105.3278
11.6703 | -105.3379
11.674 | -105.3401
11.6777 | -105.346
11.6815 | -105.3502
11.6852 | -105.3517
11.689 | -105.3516
11.6927 | -105.3522
11.6965 | -105.3491
11.7002 | -105.3441
11.7039 | -105.3411
11.7077 | -105.3343
11.7114 | -105.3271
11.7152 | -105.3152
11.7189 | -105.3032
11.7227 | -105.2914
11.7264 | -105.2804
11.7301 | -105.2716
11.7339 | -105.2655
11.7376 | -105.2571
11.7414 | -105.2473
11.7451 | -105.2388
11.7489 | -105.23
11.7526 | -105.2178
11.7563 | -105.2058
11.7601 | -105.1946
11.7638 | -105.183
11.7676 | -105.1707
11.7713 | -105.1618
11.7751 | -105.1556
11.7788 | -105.1515
11.7825 | -105.1488
11.7863 | -105.1483
11.79 | -105.1484
11.7938 | -105.1489
11.7975 | -105.151
11.8013 | -105.1557
11.805 | -105.1628
11.8087 | -105.1742
11.8125 | -105.1869
11.8162 | -105.2047
11.82 | -105.2258
11.8237 | -105.248
11.8275 | -105.2725
11.8312 | -105.2966
11.8349 | -105.3226
11.8387 | -105.3421
11.8424 | -105.368
11.8462 | -105.3908
11.8499 | -105.409
11.8537 | -105.4283
11.8574 | -105.4465
11.8611 | -105.4584
11.8649 | -105.4696
11.8686 | -105.4789
11.8724 | -105.486
11.8761 | -105.4857
11.8799 | -105.4914
11.8836 | -105.4936
11.8873 | -105.4977
11.8911 | -105.4979
11.8948 | -105.4977
11.8986 | -105.4946
11.9023 | -105.4858
11.9061 | -105.4763
11.9098 | -105.472
11.9135 | -105.4702
11.9173 | -105.4658
11.921 | -105.4616
11.9248 | -105.4576
11.9285 | -105.456
11.9323 | -105.456
11.936 | -105.4585
11.9397 | -105.4608
11.9435 | -105.4613
11.9472 | -105.4597
11.951 | -105.4593
11.9547 | -105.4611
11.9585 | -105.4625
11.9622 | -105.4637
11.9659 | -105.4707
11.9697 | -105.4785
11.9734 | -105.4848
11.9772 | -105.4876
11.9809 | -105.4903
11.9847 | -105.4901
11.9884 | -105.4866
11.9921 | -105.4812
11.9959 | -105.4756
11.9996 | -105.4699
12.0034 | -105.4627
12.0071 | -105.4543
12.0109 | -105.4446
12.0146 | -105.4268
12.0183 | -105.4077
12.0221 | -105.3912
12.0258 | -105.3754
12.0296 | -105.3576
12.0333 | -105.3396
12.0371 | -105.3216
12.0408 | -105.3055
12.0445 | -105.285
12.0483 | -105.2634
12.052 | -105.2486
12.0558 | -105.2341
12.0595 | -105.2206
12.0633 | -105.2107
12.067 | -105.2021
12.0707 | -105.1949
12.0745 | -105.1916
12.0782 | -105.1917
12.082 | -105.1921
12.0857 | -105.1913
12.0895 | -105.1904
12.0932 | -105.1908
12.0969 | -105.1939
12.1007 | -105.1982
12.1044 | -105.2042
12.1082 | -105.2103
12.1119 | -105.2163
12.1157 | -105.2224
12.1194 | -105.2326
12.1231 | -105.2456
12.1269 | -105.2561
12.1306 | -105.2654
12.1344 | -105.276
12.1381 | -105.2889
12.1419 | -105.2984
12.1456 | -105.308
12.1493 | -105.3195
12.1531 | -105.338
12.1568 | -105.3517
12.1606 | -105.3703
12.1643 | -105.3882
12.1681 | -105.4068
12.1718 | -105.4253
12.1755 | -105.4449
12.1793 | -105.4651
12.183 | -105.4839
12.1868 | -105.5017
12.1905 | -105.5214
12.1943 | -105.5425
12.198 | -105.5572
12.2017 | -105.5693
12.2055 | -105.5842
12.2092 | -105.5997
12.213 | -105.6087
12.2167 | -105.6148
12.2205 | -105.62
12.2242 | -105.6208
12.2279 | -105.6179
12.2317 | -105.6174
12.2354 | -105.6135
12.2392 | -105.6055
12.2429 | -105.5983
12.2467 | -105.588
12.2504 | -105.5764
12.2541 | -105.5653
12.2579 | -105.5545
12.2616 | -105.5439
12.2654 | -105.5363
12.2691 | -105.5301
12.2729 | -105.5204
12.2766 | -105.5097
12.2803 | -105.4995
12.2841 | -105.493
12.2878 | -105.4862
12.2916 | -105.4791
12.2953 | -105.4727
12.2991 | -105.4691
12.3028 | -105.4672
12.3065 | -105.4627
12.3103 | -105.4598
12.314 | -105.4581
12.3178 | -105.4574
12.3215 | -105.4587
12.3253 | -105.4595
12.329 | -105.4588
12.3327 | -105.4612
12.3365 | -105.4617
12.3402 | -105.4664
12.344 | -105.4716
12.3477 | -105.478
12.3515 | -105.484
12.3552 | -105.4919
12.3589 | -105.498
12.3627 | -105.5042
12.3664 | -105.5088
12.3702 | -105.5096
12.3739 | -105.5132
12.3777 | -105.5151
12.3814 | -105.5147
12.3851 | -105.5149
12.3889 | -105.5186
12.3926 | -105.5161
12.3964 | -105.515
12.4001 | -105.5119
12.4039 | -105.507
12.4076 | -105.4977
12.4113 | -105.4863
12.4151 | -105.4734
12.4188 | -105.4575
12.4226 | -105.4425
12.4263 | -105.4301
12.4301 | -105.4175
12.4338 | -105.4015
12.4375 | -105.3857
12.4413 | -105.3678
12.445 | -105.3506
12.4488 | -105.3356
12.4525 | -105.3251
12.4563 | -105.3179
12.46 | -105.3094
12.4637 | -105.2999
12.4675 | -105.2931
12.4712 | -105.2911
12.475 | -105.2912
12.4787 | -105.2908
12.4825 | -105.2936
12.4862 | -105.2982
12.4899 | -105.2996
12.4937 | -105.3022
12.4974 | -105.3073
12.5012 | -105.3161
12.5049 | -105.3274
12.5087 | -105.3404
12.5124 | -105.3523
12.5161 | -105.3649
12.5199 | -105.3768
12.5236 | -105.3874
12.5274 | -105.4007
12.5311 | -105.4143
12.5349 | -105.4255
12.5386 | -105.4406
12.5423 | -105.4539
12.5461 | -105.4644
12.5498 | -105.4755
12.5536 | -105.4847
12.5573 | -105.4909
12.5611 | -105.4966
12.5648 | -105.5025
12.5685 | -105.5048
12.5723 | -105.5072
12.576 | -105.5088
12.5798 | -105.5119
12.5835 | -105.5136
12.5873 | -105.519
12.591 | -105.5251
12.5947 | -105.5317
12.5985 | -105.5389
12.6022 | -105.5471
12.606 | -105.5586
12.6097 | -105.5707
12.6135 | -105.5817
12.6172 | -105.5966
12.6209 | -105.6125
12.6247 | -105.6298
12.6284 | -105.6471
12.6322 | -105.6616
12.6359 | -105.6786
12.6397 | -105.6956
12.6434 | -105.7131
12.6471 | -105.7314
12.6509 | -105.7509
12.6546 | -105.7678
12.6584 | -105.7805
12.6621 | -105.7951
12.6659 | -105.8091
12.6696 | -105.824
12.6733 | -105.8374
12.6771 | -105.8446
12.6808 | -105.846
12.6846 | -105.8423
12.6883 | -105.8355
12.6921 | -105.8297
12.6958 | -105.8255
12.6995 | -105.8188
12.7033 | -105.8096
12.707 | -105.7991
12.7108 | -105.7846
12.7145 | -105.7703
12.7183 | -105.7539
12.722 | -105.7406
12.7257 | -105.7296
12.7295 | -105.7172
12.7332 | -105.705
12.737 | -105.6924
12.7407 | -105.681
12.7445 | -105.6696
12.7482 | -105.6595
12.7519 | -105.6541
12.7557 | -105.6528
12.7594 | -105.6516
12.7632 | -105.6509
12.7669 | -105.6507
12.7707 | -105.6529
12.7744 | -105.6561
12.7781 | -105.6597
12.7819 | -105.6674
12.7856 | -105.676
12.7894 | -105.6822
12.7931 | -105.6919
12.7969 | -105.7059
12.8006 | -105.7197
12.8043 | -105.7325
12.8081 | -105.7434
12.8118 | -105.752
12.8156 | -105.7618
12.8193 | -105.7691
12.8231 | -105.7772
12.8268 | -105.7852
12.8305 | -105.7891
12.8343 | -105.7926
12.838 | -105.7977
12.8418 | -105.8016
12.8455 | -105.8005
12.8493 | -105.7988
12.853 | -105.7966
12.8567 | -105.7913
12.8605 | -105.7867
12.8642 | -105.7806
12.868 | -105.7722
12.8717 | -105.76
12.8755 | -105.7441
12.8792 | -105.7269
12.8829 | -105.71
12.8867 | -105.6923
12.8904 | -105.6749
12.8942 | -105.6576
12.8979 | -105.6379
12.9017 | -105.6239
12.9054 | -105.607
12.9091 | -105.5933
12.9129 | -105.5832
12.9166 | -105.5729
12.9204 | -105.5625
12.9241 | -105.5533
12.9279 | -105.5472
12.9316 | -105.5425
12.9353 | -105.539
12.9391 | -105.5374
12.9428 | -105.5374
12.9466 | -105.5365
12.9503 | -105.5376
12.9541 | -105.5344
12.9578 | -105.5336
12.9615 | -105.5351
12.9653 | -105.5364
12.969 | -105.5364
12.9728 | -105.5386
12.9765 | -105.5447
12.9803 | -105.547
12.984 | -105.5484
12.9877 | -105.5502
12.9915 | -105.5542
12.9952 | -105.5616
12.999 | -105.5695
13.0027 | -105.5774
13.0065 | -105.5871
13.0102 | -105.5989
13.0139 | -105.612
13.0177 | -105.6271
13.0214 | -105.6412
13.0252 | -105.6536
13.0289 | -105.6702
13.0327 | -105.6874
13.0364 | -105.7031
13.0401 | -105.7168
13.0439 | -105.7336
13.0476 | -105.748
13.0514 | -105.7657
13.0551 | -105.7837
13.0589 | -105.803
13.0626 | -105.8202
13.0663 | -105.8364
13.0701 | -105.8485
13.0738 | -105.8603
13.0776 | -105.8703
13.0813 | -105.8784
13.0851 | -105.8851
13.0888 | -105.884
13.0925 | -105.8825
13.0963 | -105.8788
13.1 | -105.8745
13.1038 | -105.867
13.1075 | -105.86
13.1113 | -105.8516
13.115 | -105.8386
13.1187 | -105.8253
13.1225 | -105.8116
13.1262 | -105.7982
13.13 | -105.7854
13.1337 | -105.7741
13.1375 | -105.7648
13.1412 | -105.7563
13.1449 | -105.7517
13.1487 | -105.7458
13.1524 | -105.7424
13.1562 | -105.742
13.1599 | -105.7427
13.1637 | -105.7458
13.1674 | -105.7475
13.1711 | -105.7521
13.1749 | -105.76
13.1786 | -105.7647
13.1824 | -105.7703
13.1861 | -105.7787
13.1899 | -105.7853
13.1936 | -105.7931
13.1973 | -105.7979
13.2011 | -105.8011
13.2048 | -105.8044
13.2086 | -105.8081
13.2123 | -105.8115
13.2161 | -105.8125
13.2198 | -105.8116
13.2235 | -105.8114
13.2273 | -105.81
13.231 | -105.8064
13.2348 | -105.7998
13.2385 | -105.7914
13.2423 | -105.7816
13.246 | -105.7721
13.2497 | -105.7614
13.2535 | -105.7489
13.2572 | -105.7378
13.261 | -105.7303
13.2647 | -105.7222
13.2685 | -105.7121
13.2722 | -105.7009
13.2759 | -105.6924
13.2797 | -105.6823
13.2834 | -105.6692
13.2872 | -105.6568
13.2909 | -105.6436
13.2947 | -105.6325
13.2984 | -105.6218
13.3021 | -105.6117
13.3059 | -105.6011
13.3096 | -105.5867
13.3134 | -105.5708
13.3171 | -105.5575
13.3209 | -105.5429
13.3246 | -105.5273
13.3283 | -105.5084
13.3321 | -105.4931
13.3358 | -105.4724
13.3396 | -105.4486
13.3433 | -105.4223
13.3471 | -105.3956
13.3508 | -105.3682
13.3545 | -105.3421
13.3583 | -105.3191
13.362 | -105.2957
13.3658 | -105.27
13.3695 | -105.248
13.3733 | -105.2236
13.377 | -105.1987
13.3807 | -105.1793
13.3845 | -105.1609
13.3882 | -105.1435
13.392 | -105.1283
13.3957 | -105.1167
13.3995 | -105.1102
13.4032 | -105.1071
13.4069 | -105.1084
13.4107 | -105.1097
13.4144 | -105.1146
13.4182 | -105.12
13.4219 | -105.1317
13.4257 | -105.1477
13.4294 | -105.1652
13.4331 | -105.1824
13.4369 | -105.2013
13.4406 | -105.2205
13.4444 | -105.243
13.4481 | -105.2675
13.4519 | -105.2914
13.4556 | -105.3164
13.4593 | -105.3439
13.4631 | -105.373
13.4668 | -105.4001
13.4706 | -105.4299
13.4743 | -105.4587
13.4781 | -105.4832
13.4818 | -105.51
13.4855 | -105.5363
13.4893 | -105.5581
13.493 | -105.578
13.4968 | -105.5964
13.5005 | -105.6185
13.5043 | -105.6412
13.508 | -105.6616
13.5117 | -105.6773
13.5155 | -105.6955
13.5192 | -105.7127
13.523 | -105.7266
13.5267 | -105.7366
13.5305 | -105.7495
13.5342 | -105.7642
13.5379 | -105.7778
13.5417 | -105.7911
13.5454 | -105.8055
13.5492 | -105.814
13.5529 | -105.823
13.5567 | -105.8316
13.5604 | -105.8399
13.5641 | -105.8488
13.5679 | -105.8582
13.5716 | -105.866
13.5754 | -105.8684
13.5791 | -105.8695
13.5829 | -105.8749
13.5866 | -105.8762
13.5903 | -105.8743
13.5941 | -105.8691
13.5978 | -105.8638
13.6016 | -105.8599
13.6053 | -105.854
13.6091 | -105.8447
13.6128 | -105.8367
13.6165 | -105.8216
13.6203 | -105.8037
13.624 | -105.7881
13.6278 | -105.7699
13.6315 | -105.7542
13.6353 | -105.7408
13.639 | -105.7224
13.6427 | -105.6988
13.6465 | -105.6762
13.6502 | -105.659
13.654 | -105.6401
13.6577 | -105.6212
13.6615 | -105.6055
13.6652 | -105.5906
13.6689 | -105.5789
13.6727 | -105.5709
13.6764 | -105.5611
13.6802 | -105.5544
13.6839 | -105.5504
13.6877 | -105.5467
13.6914 | -105.5413
13.6951 | -105.5393
13.6989 | -105.5407
13.7026 | -105.5454
13.7064 | -105.5511
13.7101 | -105.5571
13.7139 | -105.5629
13.7176 | -105.5724
13.7213 | -105.5866
13.7251 | -105.5988
13.7288 | -105.6083
13.7326 | -105.6186
13.7363 | -105.6324
13.7401 | -105.645
13.7438 | -105.6554
13.7475 | -105.6692
13.7513 | -105.6829
13.755 | -105.6961
13.7588 | -105.7087
13.7625 | -105.7219
13.7663 | -105.7332
13.77 | -105.7451
13.7737 | -105.7587
13.7775 | -105.7711
13.7812 | -105.7802
13.785 | -105.7923
13.7887 | -105.8015
13.7925 | -105.806
13.7962 | -105.815
13.7999 | -105.8243
13.8037 | -105.8357
13.8074 | -105.847
13.8112 | -105.8593
13.8149 | -105.8719
13.8187 | -105.8852
13.8224 | -105.8967
13.8261 | -105.9099
13.8299 | -105.9193
13.8336 | -105.9274
13.8374 | -105.9388
13.8411 | -105.9529
13.8449 | -105.966
13.8486 | -105.9746
13.8523 | -105.9866
13.8561 | -105.9973
13.8598 | -106.0017
13.8636 | -106.0065
13.8673 | -106.0112
13.8711 | -106.0174
13.8748 | -106.0253
13.8785 | -106.0296
13.8823 | -106.036
13.886 | -106.0435
13.8898 | -106.0528
13.8935 | -106.0579
13.8973 | -106.0611
13.901 | -106.0636
13.9047 | -106.0705
13.9085 | -106.0753
13.9122 | -106.0761
13.916 | -106.0743
13.9197 | -106.0738
13.9235 | -106.0768
13.9272 | -106.0769
13.9309 | -106.0778
13.9347 | -106.0746
13.9384 | -106.0711
13.9422 | -106.0706
13.9459 | -106.0703
13.9497 | -106.0664
13.9534 | -106.0625
13.9571 | -106.0612
13.9609 | -106.0547
13.9646 | -106.0469
13.9684 | -106.0389
13.9721 | -106.0283
13.9759 | -106.0185
13.9796 | -106.0097
13.9833 | -105.9989
13.9871 | -105.9905
13.9908 | -105.9836
13.9946 | -105.9704
13.9983 | -105.9593
14.0021 | -105.9459
14.0058 | -105.9356
14.0095 | -105.9258
14.0133 | -105.9176
14.017 | -105.9113
14.0208 | -105.9065
14.0245 | -105.8994
14.0283 | -105.8919
14.032 | -105.8872
14.0357 | -105.8878
14.0395 | -105.8869
14.0432 | -105.884
14.047 | -105.8814
14.0507 | -105.8795
14.0545 | -105.8764
14.0582 | -105.8788
14.0619 | -105.8858
14.0657 | -105.8919
14.0694 | -105.8956
14.0732 | -105.8998
14.0769 | -105.9032
14.0807 | -105.9077
14.0844 | -105.9109
14.0881 | -105.9146
14.0919 | -105.9189
14.0956 | -105.9253
14.0994 | -105.9287
14.1031 | -105.9314
14.1069 | -105.9311
14.1106 | -105.932
14.1143 | -105.9352
14.1181 | -105.9366
14.1218 | -105.9349
14.1256 | -105.9317
14.1293 | -105.9286
14.1331 | -105.9199
14.1368 | -105.9133
14.1405 | -105.9068
14.1443 | -105.8986
14.148 | -105.8908
14.1518 | -105.8851
14.1555 | -105.8805
14.1593 | -105.8748
14.163 | -105.8703
14.1667 | -105.8666
14.1705 | -105.8621
14.1742 | -105.8611
14.178 | -105.8589
14.1817 | -105.8563
14.1855 | -105.8574
14.1892 | -105.8576
14.1929 | -105.8603
14.1967 | -105.8643
14.2004 | -105.8664
14.2042 | -105.8734
14.2079 | -105.8805
14.2117 | -105.8866
14.2154 | -105.8943
14.2191 | -105.903
14.2229 | -105.9089
14.2266 | -105.9217
14.2304 | -105.9369
14.2341 | -105.9484
14.2379 | -105.9576
14.2416 | -105.9685
14.2453 | -105.9803
14.2491 | -105.994
14.2528 | -106.0084
14.2566 | -106.0236
14.2603 | -106.0384
14.2641 | -106.0505
14.2678 | -106.0611
14.2715 | -106.073
14.2753 | -106.0872
14.279 | -106.1012
14.2828 | -106.1179
14.2865 | -106.1318
14.2903 | -106.1467
14.294 | -106.1616
14.2977 | -106.1759
14.3015 | -106.1895
14.3052 | -106.2064
14.309 | -106.2232
14.3127 | -106.2406
14.3165 | -106.2564
14.3202 | -106.2704
14.3239 | -106.2852
14.3277 | -106.2963
14.3314 | -106.307
14.3352 | -106.3197
14.3389 | -106.3311
14.3427 | -106.3376
14.3464 | -106.3441
14.3501 | -106.3501
14.3539 | -106.3528
14.3576 | -106.3545
14.3614 | -106.3554
14.3651 | -106.3515
14.3689 | -106.349
14.3726 | -106.3458
14.3763 | -106.3404
14.3801 | -106.3331
14.3838 | -106.3215
14.3876 | -106.3087
14.3913 | -106.2958
14.3951 | -106.2851
14.3988 | -106.2712
14.4025 | -106.2552
14.4063 | -106.2397
14.41 | -106.2252
14.4138 | -106.2113
14.4175 | -106.1988
14.4213 | -106.1863
14.425 | -106.1719
14.4287 | -106.1603
14.4325 | -106.1442
14.4362 | -106.1283
14.44 | -106.1127
14.4437 | -106.1009
14.4475 | -106.0861
14.4512 | -106.0704
14.4549 | -106.0586
14.4587 | -106.044
14.4624 | -106.0281
14.4662 | -106.0128
14.4699 | -105.9995
14.4737 | -105.9916
14.4774 | -105.9846
14.4811 | -105.9739
14.4849 | -105.964
14.4886 | -105.9535
14.4924 | -105.9481
14.4961 | -105.9414
14.4999 | -105.9334
14.5036 | -105.9249
14.5073 | -105.9189
14.5111 | -105.913
14.5148 | -105.9065
14.5186 | -105.9007
14.5223 | -105.8987
14.5261 | -105.8983
14.5298 | -105.8959
14.5335 | -105.897
14.5373 | -105.9012
14.541 | -105.9036
14.5448 | -105.9096
14.5485 | -105.9199
14.5523 | -105.9291
14.556 | -105.9358
14.5597 | -105.9445
14.5635 | -105.9546
14.5672 | -105.9634
14.571 | -105.9766
14.5747 | -105.9911
14.5785 | -106.0038
14.5822 | -106.016
14.5859 | -106.0297
14.5897 | -106.0383
14.5934 | -106.0486
14.5972 | -106.0605
14.6009 | -106.0688
14.6047 | -106.0754
14.6084 | -106.0824
14.6121 | -106.0893
14.6159 | -106.0941
14.6196 | -106.0991
14.6234 | -106.1008
14.6271 | -106.101
14.6309 | -106.1031
14.6346 | -106.1059
14.6383 | -106.1088
14.6421 | -106.1131
14.6458 | -106.1168
14.6496 | -106.1222
14.6533 | -106.1262
14.6571 | -106.1296
14.6608 | -106.1339
14.6645 | -106.135
14.6683 | -106.1384
14.672 | -106.142
14.6758 | -106.1481
14.6795 | -106.1528
14.6833 | -106.155
14.687 | -106.1597
14.6907 | -106.1677
14.6945 | -106.1754
14.6982 | -106.1855
14.702 | -106.1944
14.7057 | -106.2043
14.7095 | -106.2172
14.7132 | -106.2299
14.7169 | -106.2415
14.7207 | -106.2514
14.7244 | -106.2625
14.7282 | -106.2757
14.7319 | -106.2876
14.7357 | -106.2979
14.7394 | -106.3091
14.7431 | -106.3176
14.7469 | -106.3236
14.7506 | -106.3295
14.7544 | -106.3322
14.7581 | -106.3324
14.7619 | -106.331
14.7656 | -106.3311
14.7693 | -106.3298
14.7731 | -106.3245
14.7768 | -106.3196
14.7806 | -106.3096
14.7843 | -106.2955
14.7881 | -106.2831
14.7918 | -106.2692
14.7955 | -106.253
14.7993 | -106.2342
14.803 | -106.2192
14.8068 | -106.2018
14.8105 | -106.1797
14.8143 | -106.1561
14.818 | -106.1346
14.8217 | -106.1132
14.8255 | -106.0919
14.8292 | -106.0702
14.833 | -106.0457
14.8367 | -106.0246
14.8405 | -106.0075
14.8442 | -105.9871
14.8479 | -105.9698
14.8517 | -105.9569
14.8554 | -105.9437
14.8592 | -105.9337
14.8629 | -105.9277
14.8667 | -105.9222
14.8704 | -105.9206
14.8741 | -105.922
14.8779 | -105.9279
14.8816 | -105.9368
14.8854 | -105.9472
14.8891 | -105.9543
14.8929 | -105.9638
14.8966 | -105.9781
14.9003 | -105.9927
14.9041 | -106.0081
14.9078 | -106.0239
14.9116 | -106.0407
14.9153 | -106.056
14.9191 | -106.0691
14.9228 | -106.0799
14.9265 | -106.0931
14.9303 | -106.1066
14.934 | -106.1191
14.9378 | -106.1307
14.9415 | -106.1378
14.9453 | -106.1426
14.949 | -106.144
14.9527 | -106.1476
14.9565 | -106.1532
14.9602 | -106.1557
14.964 | -106.1587
14.9677 | -106.1613
14.9715 | -106.1596
14.9752 | -106.1569
14.9789 | -106.1522
14.9827 | -106.145
14.9864 | -106.1377
14.9902 | -106.1273
14.9939 | -106.1186
14.9977 | -106.1117
15.0014 | -106.1035
15.0051 | -106.0969
15.0089 | -106.0889
15.0126 | -106.0797
15.0164 | -106.0716
15.0201 | -106.0659
15.0239 | -106.0602
15.0276 | -106.0561
15.0313 | -106.0504
15.0351 | -106.0455
15.0388 | -106.043
15.0426 | -106.0436
15.0463 | -106.044
15.0501 | -106.0436
15.0538 | -106.0434
15.0575 | -106.0487
15.0613 | -106.0539
15.065 | -106.0585
15.0688 | -106.064
15.0725 | -106.0717
15.0763 | -106.0809
15.08 | -106.09
15.0837 | -106.0989
15.0875 | -106.1089
15.0912 | -106.118
15.095 | -106.1295
15.0987 | -106.142
15.1025 | -106.156
15.1062 | -106.1664
15.1099 | -106.1746
15.1137 | -106.1821
15.1174 | -106.1888
15.1212 | -106.1955
15.1249 | -106.1997
15.1287 | -106.2009
15.1324 | -106.2037
15.1361 | -106.2075
15.1399 | -106.211
15.1436 | -106.2106
15.1474 | -106.2068
15.1511 | -106.2048
15.1549 | -106.2004
15.1586 | -106.1945
15.1623 | -106.1898
15.1661 | -106.1865
15.1698 | -106.1784
15.1736 | -106.1683
15.1773 | -106.1582
15.1811 | -106.1486
15.1848 | -106.1396
15.1885 | -106.1321
15.1923 | -106.1249
15.196 | -106.1166
15.1998 | -106.1113
15.2035 | -106.1049
15.2073 | -106.0992
15.211 | -106.0948
15.2147 | -106.0879
15.2185 | -106.0815
15.2222 | -106.0763
15.226 | -106.0723
15.2297 | -106.0709
15.2335 | -106.0707
15.2372 | -106.0709
15.2409 | -106.0727
15.2447 | -106.0756
15.2484 | -106.0779
15.2522 | -106.0832
15.2559 | -106.0864
15.2597 | -106.0875
15.2634 | -106.0923
15.2671 | -106.098
15.2709 | -106.103
15.2746 | -106.1066
15.2784 | -106.1107
15.2821 | -106.1175
15.2859 | -106.1261
15.2896 | -106.1312
15.2933 | -106.1335
15.2971 | -106.1415
15.3008 | -106.15
15.3046 | -106.1556
15.3083 | -106.1629
15.3121 | -106.1677
15.3158 | -106.1704
15.3195 | -106.1766
15.3233 | -106.1843
15.327 | -106.1888
15.3308 | -106.1936
15.3345 | -106.1962
15.3383 | -106.1999
15.342 | -106.2048
15.3457 | -106.2079
15.3495 | -106.2095
15.3532 | -106.212
15.357 | -106.2153
15.3607 | -106.2166
15.3645 | -106.217
15.3682 | -106.2203
15.3719 | -106.2222
15.3757 | -106.2219
15.3794 | -106.2244
15.3832 | -106.2268
15.3869 | -106.2244
15.3907 | -106.2246
15.3944 | -106.2243
15.3981 | -106.2217
15.4019 | -106.2197
15.4056 | -106.2161
15.4094 | -106.214
15.4131 | -106.2126
15.4169 | -106.2126
15.4206 | -106.2105
15.4243 | -106.2084
15.4281 | -106.2066
15.4318 | -106.204
15.4356 | -106.2042
15.4393 | -106.2029
15.4431 | -106.2015
15.4468 | -106.2014
15.4505 | -106.2003
15.4543 | -106.1981
15.458 | -106.196
15.4618 | -106.1935
15.4655 | -106.1922
15.4693 | -106.1925
15.473 | -106.1926
15.4767 | -106.1965
15.4805 | -106.2013
15.4842 | -106.205
15.488 | -106.2106
15.4917 | -106.2169
15.4955 | -106.2239
15.4992 | -106.2294
15.5029 | -106.2353
15.5067 | -106.2408
15.5104 | -106.2493
15.5142 | -106.2605
15.5179 | -106.2694
15.5217 | -106.2796
15.5254 | -106.2895
15.5291 | -106.2973
15.5329 | -106.3049
15.5366 | -106.3176
15.5404 | -106.3276
15.5441 | -106.3337
15.5479 | -106.3421
15.5516 | -106.3515
15.5553 | -106.3625
15.5591 | -106.3709
15.5628 | -106.3797
15.5666 | -106.3908
15.5703 | -106.4024
15.5741 | -106.4093
15.5778 | -106.4144
15.5815 | -106.4207
15.5853 | -106.4238
15.589 | -106.4243
15.5928 | -106.4265
15.5965 | -106.4304
15.6003 | -106.4301
15.604 | -106.4252
15.6077 | -106.4249
15.6115 | -106.4202
15.6152 | -106.4162
15.619 | -106.4136
15.6227 | -106.4083
15.6265 | -106.3993
15.6302 | -106.3891
15.6339 | -106.3792
15.6377 | -106.3668
15.6414 | -106.3556
15.6452 | -106.3445
15.6489 | -106.33
15.6527 | -106.3118
15.6564 | -106.2966
15.6601 | -106.2805
15.6639 | -106.2671
15.6676 | -106.2539
15.6714 | -106.2405
15.6751 | -106.2272
15.6789 | -106.2128
15.6826 | -106.1991
15.6863 | -106.1865
15.6901 | -106.1744
15.6938 | -106.1648
15.6976 | -106.1539
15.7013 | -106.1428
15.7051 | -106.134
15.7088 | -106.124
15.7125 | -106.1125
15.7163 | -106.1021
15.72 | -106.0918
15.7238 | -106.0849
15.7275 | -106.0794
15.7313 | -106.0746
15.735 | -106.0712
15.7387 | -106.0678
15.7425 | -106.0636
15.7462 | -106.0618
15.75 | -106.0634
15.7537 | -106.0654
15.7575 | -106.0685
15.7612 | -106.0696
15.7649 | -106.0665
15.7687 | -106.068
15.7724 | -106.0684
15.7762 | -106.0667
15.7799 | -106.065
15.7837 | -106.0678
15.7874 | -106.0705
15.7911 | -106.0706
15.7949 | -106.0707
15.7986 | -106.069
15.8024 | -106.0672
15.8061 | -106.0651
15.8099 | -106.0651
15.8136 | -106.064
15.8173 | -106.0607
15.8211 | -106.058
15.8248 | -106.0573
15.8286 | -106.0583
15.8323 | -106.0573
15.8361 | -106.0563
15.8398 | -106.0533
15.8435 | -106.0524
15.8473 | -106.0524
15.851 | -106.0551
15.8548 | -106.0578
15.8585 | -106.0601
15.8623 | -106.0654
15.866 | -106.0725
15.8697 | -106.083
15.8735 | -106.096
15.8772 | -106.1104
15.881 | -106.1224
15.8847 | -106.1369
15.8885 | -106.1513
15.8922 | -106.1657
15.8959 | -106.1819
15.8997 | -106.1988
15.9034 | -106.2173
15.9072 | -106.2351
15.9109 | -106.2533
15.9147 | -106.2716
15.9184 | -106.2879
15.9221 | -106.3018
15.9259 | -106.3163
15.9296 | -106.3315
15.9334 | -106.3458
15.9371 | -106.362
15.9409 | -106.3759
15.9446 | -106.3876
15.9483 | -106.3975
15.9521 | -106.4014
15.9558 | -106.4023
15.9596 | -106.4057
15.9633 | -106.4062
15.9671 | -106.4043
15.9708 | -106.4031
15.9745 | -106.3976
15.9783 | -106.3914
15.982 | -106.3859
15.9858 | -106.3764
15.9895 | -106.3679
15.9933 | -106.3636
15.997 | -106.3558
16.0007 | -106.3479
16.0045 | -106.3407
16.0082 | -106.3315
16.012 | -106.3208
16.0157 | -106.3111
16.0195 | -106.306
16.0232 | -106.3033
16.0269 | -106.3011
16.0307 | -106.3002
16.0344 | -106.2995
16.0382 | -106.3007
16.0419 | -106.3042
16.0457 | -106.3082
16.0494 | -106.3122
16.0531 | -106.3153
16.0569 | -106.321
16.0606 | -106.3285
16.0644 | -106.3386
16.0681 | -106.3488
16.0719 | -106.3549
16.0756 | -106.3594
16.0793 | -106.3668
16.0831 | -106.3717
16.0868 | -106.3739
16.0906 | -106.3778
16.0943 | -106.3805
16.0981 | -106.3802
16.1018 | -106.3784
16.1055 | -106.3759
16.1093 | -106.3708
16.113 | -106.3625
16.1168 | -106.3555
16.1205 | -106.3465
16.1243 | -106.3353
16.128 | -106.3227
16.1317 | -106.308
16.1355 | -106.2936
16.1392 | -106.2771
16.143 | -106.2584
16.1467 | -106.2385
16.1505 | -106.2205
16.1542 | -106.2035
16.1579 | -106.1855
16.1617 | -106.1689
16.1654 | -106.1518
16.1692 | -106.1356
16.1729 | -106.1225
16.1767 | -106.1117
16.1804 | -106.1046
16.1841 | -106.0989
16.1879 | -106.0909
16.1916 | -106.0871
16.1954 | -106.0869
16.1991 | -106.1015
16.2029 | -106.1022
16.2066 | -106.105
16.2103 | -106.1117
16.2141 | -106.1183
16.2178 | -106.1287
16.2216 | -106.1382
16.2253 | -106.1487
16.2291 | -106.1589
16.2328 | -106.1689
16.2365 | -106.1802
16.2403 | -106.1921
16.244 | -106.2032
16.2478 | -106.2155
16.2515 | -106.227
16.2553 | -106.2372
16.259 | -106.2494
16.2627 | -106.2579
16.2665 | -106.2659
16.2702 | -106.2722
16.274 | -106.2795
16.2777 | -106.2847
16.2815 | -106.2882
16.2852 | -106.2893
16.2889 | -106.2891
16.2927 | -106.2901
16.2964 | -106.29
16.3002 | -106.29
16.3039 | -106.2915
16.3077 | -106.2921
16.3114 | -106.29
16.3151 | -106.2881
16.3189 | -106.2876
16.3226 | -106.2836
16.3264 | -106.2795
16.3301 | -106.2765
16.3339 | -106.2748
16.3376 | -106.2749
16.3413 | -106.2752
16.3451 | -106.2755
16.3488 | -106.2738
16.3526 | -106.2729
16.3563 | -106.2713
16.3601 | -106.2691
16.3638 | -106.2662
16.3675 | -106.2641
16.3713 | -106.2606
16.375 | -106.2574
16.3788 | -106.2567
16.3825 | -106.2546
16.3863 | -106.2504
16.39 | -106.2464
16.3937 | -106.2423
16.3975 | -106.2408
16.4012 | -106.2399
16.405 | -106.2381
16.4087 | -106.2381
16.4125 | -106.2339
16.4162 | -106.2315
16.4199 | -106.2305
16.4237 | -106.2263
16.4274 | -106.2223
16.4312 | -106.2184
16.4349 | -106.2147
16.4387 | -106.2145
16.4424 | -106.2158
16.4461 | -106.2146
16.4499 | -106.2149
16.4536 | -106.2155
16.4574 | -106.2165
16.4611 | -106.2173
16.4649 | -106.22
16.4686 | -106.2216
16.4723 | -106.2244
16.4761 | -106.2281
16.4798 | -106.2319
16.4836 | -106.2344
16.4873 | -106.2356
16.4911 | -106.2376
16.4948 | -106.242
16.4986 | -106.2484
16.5023 | -106.2531
16.506 | -106.2581
16.5098 | -106.2607
16.5135 | -106.2676
16.5173 | -106.2752
16.521 | -106.2832
16.5248 | -106.2904
16.5285 | -106.2945
16.5322 | -106.3
16.536 | -106.3054
16.5397 | -106.3122
16.5435 | -106.3179
16.5472 | -106.3257
16.551 | -106.3334
16.5547 | -106.3394
16.5584 | -106.3455
16.5622 | -106.3524
16.5659 | -106.3595
16.5697 | -106.3673
16.5734 | -106.3733
16.5771 | -106.3802
16.5809 | -106.3874
16.5846 | -106.3936
16.5884 | -106.3982
16.5921 | -106.4037
16.5959 | -106.4114
16.5996 | -106.4175
16.6033 | -106.4215
16.6071 | -106.4279
16.6108 | -106.4348
16.6146 | -106.4406
16.6183 | -106.4454
16.6221 | -106.4494
16.6258 | -106.4533
16.6295 | -106.4573
16.6333 | -106.4598
16.637 | -106.4638
16.6408 | -106.4678
16.6445 | -106.4705
16.6483 | -106.4775
16.652 | -106.4813
16.6558 | -106.487
16.6595 | -106.4942
16.6632 | -106.5013
16.667 | -106.5081
16.6707 | -106.5124
16.6745 | -106.5147
16.6782 | -106.5161
16.682 | -106.5186
16.6857 | -106.5202
16.6894 | -106.5233
16.6932 | -106.5281
16.6969 | -106.5324
16.7007 | -106.5353
16.7044 | -106.5406
16.7081 | -106.5435
16.7119 | -106.5452
16.7156 | -106.547
16.7194 | -106.5483
16.7231 | -106.5512
16.7269 | -106.5555
16.7306 | -106.5572
16.7343 | -106.559
16.7381 | -106.564
16.7418 | -106.5689
16.7456 | -106.5723
16.7493 | -106.5763
16.7531 | -106.58
16.7568 | -106.5863
16.7605 | -106.5921
16.7643 | -106.5934
16.768 | -106.5951
16.7718 | -106.5967
16.7755 | -106.6016
16.7793 | -106.606
16.783 | -106.6112
16.7868 | -106.6149
16.7905 | -106.6198
16.7942 | -106.6248
16.798 | -106.6279
16.8017 | -106.6303
16.8055 | -106.6328
16.8092 | -106.6352
16.813 | -106.6381
16.8167 | -106.6408
16.8204 | -106.645
16.8242 | -106.651
16.8279 | -106.6555
16.8317 | -106.6559
16.8354 | -106.6559
16.8392 | -106.6589
16.8429 | -106.6581
16.8466 | -106.6552
16.8504 | -106.6496
16.8541 | -106.6451
16.8579 | -106.6407
16.8616 | -106.6351
16.8653 | -106.6291
16.8691 | -106.6248
16.8728 | -106.6204
16.8766 | -106.6145
16.8803 | -106.608
16.8841 | -106.6026
16.8878 | -106.5986
16.8915 | -106.596
16.8953 | -106.5914
16.899 | -106.5862
16.9028 | -106.5834
16.9065 | -106.5816
16.9103 | -106.5804
16.914 | -106.5803
16.9178 | -106.5802
16.9215 | -106.58
16.9252 | -106.5809
16.929 | -106.5801
16.9327 | -106.5787
16.9365 | -106.5777
16.9402 | -106.5789
16.944 | -106.579
16.9477 | -106.5796
16.9514 | -106.581
16.9552 | -106.5822
16.9589 | -106.5831
16.9627 | -106.5846
16.9664 | -106.5838
16.9702 | -106.5852
16.9739 | -106.5844
16.9776 | -106.5819
16.9814 | -106.5802
16.9851 | -106.5761
16.9889 | -106.5732
16.9926 | -106.5693
16.9964 | -106.5628
17.0001 | -106.5557
17.0038 | -106.5491
17.0076 | -106.5459
17.0113 | -106.5437
17.0151 | -106.5399
17.0188 | -106.5324
17.0226 | -106.5216
17.0263 | -106.5123
17.03 | -106.501
17.0338 | -106.4902
17.0375 | -106.4822
17.0413 | -106.4746
17.045 | -106.4656
17.0488 | -106.456
17.0525 | -106.4471
17.0562 | -106.4389
17.06 | -106.4285
17.0637 | -106.4187
17.0675 | -106.4106
17.0712 | -106.4051
17.075 | -106.4029
17.0787 | -106.3991
17.0824 | -106.3957
17.0862 | -106.3972
17.0899 | -106.3983
17.0937 | -106.4014
17.0974 | -106.4045
17.1012 | -106.4087
17.1049 | -106.4149
17.1086 | -106.4211
17.1124 | -106.4247
17.1161 | -106.4309
17.1199 | -106.441
17.1236 | -106.4502
17.1274 | -106.4597
17.1311 | -106.4709
17.1348 | -106.4838
17.1386 | -106.4986
17.1423 | -106.5132
17.1461 | -106.5256
17.1498 | -106.5395
17.1536 | -106.5551
17.1573 | -106.5733
17.161 | -106.5893
17.1648 | -106.6021
17.1685 | -106.617
17.1723 | -106.6318
17.176 | -106.6457
17.1798 | -106.6591
17.1835 | -106.6701
17.1872 | -106.6829
17.191 | -106.6942
17.1947 | -106.7032
17.1985 | -106.7121
17.2022 | -106.7194
17.206 | -106.7241
17.2097 | -106.7274
17.2134 | -106.7314
17.2172 | -106.7357
17.2209 | -106.7381
17.2247 | -106.7408
17.2284 | -106.7433
17.2322 | -106.7437
17.2359 | -106.7409
17.2396 | -106.7397
17.2434 | -106.7346
17.2471 | -106.7292
17.2509 | -106.7239
17.2546 | -106.7161
17.2584 | -106.7064
17.2621 | -106.6982
17.2658 | -106.6918
17.2696 | -106.6837
17.2733 | -106.6763
17.2771 | -106.6645
17.2808 | -106.6531
17.2846 | -106.6379
17.2883 | -106.6259
17.292 | -106.6151
17.2958 | -106.6035
17.2995 | -106.5911
17.3033 | -106.5769
17.307 | -106.5637
17.3108 | -106.5498
17.3145 | -106.5378
17.3182 | -106.5254
17.322 | -106.5141
17.3257 | -106.5026
17.3295 | -106.4926
17.3332 | -106.483
17.337 | -106.4737
17.3407 | -106.4639
17.3444 | -106.4549
17.3482 | -106.448
17.3519 | -106.4428
17.3557 | -106.4393
17.3594 | -106.4371
17.3632 | -106.434
17.3669 | -106.4307
17.3706 | -106.4279
17.3744 | -106.4247
17.3781 | -106.42
17.3819 | -106.4145
17.3856 | -106.4117
17.3894 | -106.4115
17.3931 | -106.4136
17.3968 | -106.4158
17.4006 | -106.4162
17.4043 | -106.4162
17.4081 | -106.419
17.4118 | -106.4226
17.4156 | -106.4249
17.4193 | -106.4283
17.423 | -106.4306
17.4268 | -106.4326
17.4305 | -106.4343
17.4343 | -106.4355
17.438 | -106.4367
17.4418 | -106.4383
17.4455 | -106.4403
17.4492 | -106.4428
17.453 | -106.4464
17.4567 | -106.4513
17.4605 | -106.4545
17.4642 | -106.4577
17.468 | -106.4613
17.4717 | -106.4638
17.4754 | -106.4653
17.4792 | -106.4672
17.4829 | -106.4705
17.4867 | -106.4699
17.4904 | -106.4739
17.4942 | -106.4784
17.4979 | -106.4798
17.5016 | -106.4807
17.5054 | -106.4821
17.5091 | -106.4836
17.5129 | -106.4832
17.5166 | -106.4844
17.5204 | -106.4851
17.5241 | -106.485
17.5278 | -106.4864
17.5316 | -106.4906
17.5353 | -106.4941
17.5391 | -106.4947
17.5428 | -106.4949
17.5466 | -106.4954
17.5503 | -106.4974
17.554 | -106.5011
17.5578 | -106.5053
17.5615 | -106.5083
17.5653 | -106.5084
17.569 | -106.511
17.5728 | -106.5115
17.5765 | -106.5132
17.5802 | -106.5136
17.584 | -106.5145
17.5877 | -106.5163
17.5915 | -106.5184
17.5952 | -106.5197
17.599 | -106.521
17.6027 | -106.5239
17.6064 | -106.526
17.6102 | -106.5289
17.6139 | -106.5316
17.6177 | -106.5338
17.6214 | -106.5399
17.6252 | -106.5449
17.6289 | -106.5487
17.6326 | -106.5533
17.6364 | -106.5549
17.6401 | -106.557
17.6439 | -106.558
17.6476 | -106.5586
17.6514 | -106.5613
17.6551 | -106.5644
17.6588 | -106.5669
17.6626 | -106.5704
17.6663 | -106.5734
17.6701 | -106.5776
17.6738 | -106.585
17.6776 | -106.5898
17.6813 | -106.5936
17.685 | -106.5966
17.6888 | -106.6007
17.6925 | -106.6079
17.6963 | -106.6136
17.7 | -106.6189
17.7038 | -106.6234
17.7075 | -106.6261
17.7112 | -106.6272
17.715 | -106.6304
17.7187 | -106.636
17.7225 | -106.6404
17.7262 | -106.6471
17.73 | -106.6525
17.7337 | -106.6554
17.7374 | -106.6596
17.7412 | -106.663
17.7449 | -106.6653
17.7487 | -106.6686
17.7524 | -106.673
17.7562 | -106.6778
17.7599 | -106.6812
17.7636 | -106.6859
17.7674 | -106.6895
17.7711 | -106.6937
17.7749 | -106.6976
17.7786 | -106.7029
17.7824 | -106.7065
17.7861 | -106.7109
17.7898 | -106.7141
17.7936 | -106.7174
17.7973 | -106.7216
17.8011 | -106.7245
17.8048 | -106.7292
17.8086 | -106.7318
17.8123 | -106.7323
17.816 | -106.7292
17.8198 | -106.7339
17.8235 | -106.7344
17.8273 | -106.7357
17.831 | -106.7357
17.8348 | -106.7345
17.8385 | -106.7328
17.8422 | -106.7289
17.846 | -106.725
17.8497 | -106.7193
17.8535 | -106.7122
17.8572 | -106.7036
17.861 | -106.6956
17.8647 | -106.6847
17.8684 | -106.6745
17.8722 | -106.6657
17.8759 | -106.6538
17.8797 | -106.6413
17.8834 | -106.6286
17.8872 | -106.6159
17.8909 | -106.6048
17.8946 | -106.5913
17.8984 | -106.5767
17.9021 | -106.5645
17.9059 | -106.5524
17.9096 | -106.5396
17.9134 | -106.5271
17.9171 | -106.5153
17.9208 | -106.5043
17.9246 | -106.4922
17.9283 | -106.4786
17.9321 | -106.4658
17.9358 | -106.4551
17.9396 | -106.4469
17.9433 | -106.4385
17.947 | -106.4276
17.9508 | -106.4191
17.9545 | -106.4114
17.9583 | -106.4046
17.962 | -106.3968
17.9658 | -106.3899
17.9695 | -106.3835
17.9732 | -106.3757
17.977 | -106.3698
17.9807 | -106.3625
17.9845 | -106.3528
17.9882 | -106.3442
17.992 | -106.3388
17.9957 | -106.3311
17.9994 | -106.3213
18.0032 | -106.3122
18.0069 | -106.3038
18.0107 | -106.2958
18.0144 | -106.2877
18.0182 | -106.2806
18.0219 | -106.2709
18.0256 | -106.2617
18.0294 | -106.2549
18.0331 | -106.2477
18.0369 | -106.2414
18.0406 | -106.2357
18.0444 | -106.2307
18.0481 | -106.2234
18.0518 | -106.219
18.0556 | -106.214
18.0593 | -106.2076
18.0631 | -106.2048
18.0668 | -106.2042
18.0706 | -106.2028
18.0743 | -106.1972
18.078 | -106.1907
18.0818 | -106.1871
18.0855 | -106.1818
18.0893 | -106.1776
18.093 | -106.1732
18.0968 | -106.169
18.1005 | -106.1683
18.1042 | -106.1666
18.108 | -106.165
18.1117 | -106.163
18.1155 | -106.16
18.1192 | -106.1558
18.123 | -106.1524
18.1267 | -106.1479
18.1304 | -106.1427
18.1342 | -106.139
18.1379 | -106.1374
18.1417 | -106.137
18.1454 | -106.1351
18.1492 | -106.1309
18.1529 | -106.1266
18.1566 | -106.1244
18.1604 | -106.1205
18.1641 | -106.1151
18.1679 | -106.1122
18.1716 | -106.1094
18.1754 | -106.1068
18.1791 | -106.1055
18.1828 | -106.1028
18.1866 | -106.102
18.1903 | -106.1029
18.1941 | -106.1008
18.1978 | -106.0955
18.2016 | -106.0896
18.2053 | -106.086
18.209 | -106.0823
18.2128 | -106.0793
18.2165 | -106.0777
18.2203 | -106.076
18.224 | -106.074
18.2278 | -106.0707
18.2315 | -106.0682
18.2352 | -106.0652
18.239 | -106.063
18.2427 | -106.064
18.2465 | -106.0636
18.2502 | -106.0617
18.254 | -106.0621
18.2577 | -106.0634
18.2614 | -106.0638
18.2652 | -106.063
18.2689 | -106.0644
18.2727 | -106.0667
18.2764 | -106.0683
18.2802 | -106.0715
18.2839 | -106.0754
18.2876 | -106.0771
18.2914 | -106.0773
18.2951 | -106.079
18.2989 | -106.0822
18.3026 | -106.0858
18.3064 | -106.0916
18.3101 | -106.0933
18.3138 | -106.095
18.3176 | -106.0966
18.3213 | -106.0979
18.3251 | -106.1005
18.3288 | -106.1044
18.3326 | -106.1103
18.3363 | -106.1144
18.34 | -106.1195
18.3438 | -106.1261
18.3475 | -106.1323
18.3513 | -106.1387
18.355 | -106.145
18.3588 | -106.1503
18.3625 | -106.1549
18.3662 | -106.1584
18.37 | -106.1614
18.3737 | -106.1661
18.3775 | -106.173
18.3812 | -106.1792
18.385 | -106.1843
18.3887 | -106.1903
18.3924 | -106.1965
18.3962 | -106.2024
18.3999 | -106.2064
18.4037 | -106.2105
18.4074 | -106.2147
18.4112 | -106.2204
18.4149 | -106.227
18.4186 | -106.232
18.4224 | -106.2367
18.4261 | -106.242
18.4299 | -106.245
18.4336 | -106.2492
18.4374 | -106.2533
18.4411 | -106.2552
18.4448 | -106.2613
18.4486 | -106.2676
18.4523 | -106.2745
18.4561 | -106.2818
18.4598 | -106.2882
18.4636 | -106.2953
18.4673 | -106.3017
18.471 | -106.3075
18.4748 | -106.3153
18.4785 | -106.3211
18.4823 | -106.3307
18.486 | -106.3404
18.4898 | -106.3507
18.4935 | -106.3604
18.4972 | -106.3709
18.501 | -106.3814
18.5047 | -106.3905
18.5085 | -106.3992
18.5122 | -106.4066
18.516 | -106.4117
18.5197 | -106.42
18.5234 | -106.4285
18.5272 | -106.4353
18.5309 | -106.4402
18.5347 | -106.4464
18.5384 | -106.4512
18.5422 | -106.4571
18.5459 | -106.4639
18.5496 | -106.4715
18.5534 | -106.4796
18.5571 | -106.484
18.5609 | -106.4871
18.5646 | -106.4928
18.5684 | -106.4976
18.5721 | -106.5035
18.5758 | -106.5088
18.5796 | -106.5156
18.5833 | -106.5216
18.5871 | -106.5263
18.5908 | -106.5316
18.5946 | -106.5353
18.5983 | -106.5406
18.602 | -106.5439
18.6058 | -106.549
18.6095 | -106.5545
18.6133 | -106.5604
18.617 | -106.566
18.6208 | -106.5736
18.6245 | -106.5821
18.6282 | -106.588
18.632 | -106.5964
18.6357 | -106.6051
18.6395 | -106.6133
18.6432 | -106.6195
18.647 | -106.626
18.6507 | -106.6351
18.6544 | -106.6449
18.6582 | -106.6542
18.6619 | -106.6637
18.6657 | -106.6744
18.6694 | -106.6848
18.6732 | -106.6975
18.6769 | -106.7117
18.6806 | -106.7238
18.6844 | -106.7357
18.6881 | -106.7499
18.6919 | -106.7662
18.6956 | -106.7818
18.6994 | -106.7955
18.7031 | -106.8102
18.7068 | -106.8273
18.7106 | -106.8415
18.7143 | -106.8539
18.7181 | -106.8663
18.7218 | -106.8799
18.7256 | -106.8963
18.7293 | -106.9107
18.733 | -106.9224
18.7368 | -106.9337
18.7405 | -106.9424
18.7443 | -106.952
18.748 | -106.9618
18.7518 | -106.9695
18.7555 | -106.9748
18.7592 | -106.979
18.763 | -106.9791
18.7667 | -106.9807
18.7705 | -106.9839
18.7742 | -106.9859
18.778 | -106.9837
18.7817 | -106.983
18.7854 | -106.9791
18.7892 | -106.9761
18.7929 | -106.9746
18.7967 | -106.9692
18.8004 | -106.9603
18.8042 | -106.9533
18.8079 | -106.9441
18.8116 | -106.9325
18.8154 | -106.9221
18.8191 | -106.9106
18.8229 | -106.9019
18.8266 | -106.8916
18.8304 | -106.8784
18.8341 | -106.8661
18.8378 | -106.8531
18.8416 | -106.8401
18.8453 | -106.8266
18.8491 | -106.8128
18.8528 | -106.8002
18.8566 | -106.7857
18.8603 | -106.7758
18.864 | -106.7667
18.8678 | -106.7593
18.8715 | -106.7497
18.8753 | -106.741
18.879 | -106.7332
18.8828 | -106.7263
18.8865 | -106.7202
18.8902 | -106.7143
18.894 | -106.712
18.8977 | -106.7166
18.9015 | -106.7159
18.9052 | -106.7173
18.909 | -106.7184
18.9127 | -106.7211
18.9164 | -106.7243
18.9202 | -106.7279
18.9239 | -106.7307
18.9277 | -106.7327
18.9314 | -106.7364
18.9352 | -106.7414
18.9389 | -106.7471
18.9426 | -106.7511
18.9464 | -106.7545
18.9501 | -106.762
18.9539 | -106.7679
18.9576 | -106.7733
18.9614 | -106.7768
18.9651 | -106.7811
18.9688 | -106.7848
18.9726 | -106.7864
18.9763 | -106.7871
18.9801 | -106.788
18.9838 | -106.7893
18.9876 | -106.7893
18.9913 | -106.7887
18.995 | -106.7854
18.9988 | -106.7849
19.0025 | -106.7855
19.0063 | -106.7845
19.01 | -106.7819
19.0138 | -106.7787
19.0175 | -106.7753
19.0212 | -106.7731
19.025 | -106.7692
19.0287 | -106.7678
19.0325 | -106.7681
19.0362 | -106.7657
19.04 | -106.7616
19.0437 | -106.7594
19.0474 | -106.758
19.0512 | -106.7558
19.0549 | -106.7552
19.0587 | -106.7549
19.0624 | -106.7554
19.0662 | -106.7572
19.0699 | -106.7606
19.0736 | -106.7631
19.0774 | -106.7636
19.0811 | -106.7637
19.0849 | -106.7637
19.0886 | -106.7649
19.0924 | -106.7657
19.0961 | -106.7686
19.0998 | -106.772
19.1036 | -106.7746
19.1073 | -106.7797
19.1111 | -106.7849
19.1148 | -106.7868
19.1186 | -106.7864
19.1223 | -106.7883
19.126 | -106.792
19.1298 | -106.792
19.1335 | -106.79
19.1373 | -106.7889
19.141 | -106.7898
19.1448 | -106.7891
19.1485 | -106.7861
19.1522 | -106.7806
19.156 | -106.7754
19.1597 | -106.7705
19.1635 | -106.7657
19.1672 | -106.7608
19.171 | -106.754
19.1747 | -106.745
19.1784 | -106.7345
19.1822 | -106.7256
19.1859 | -106.7161
19.1897 | -106.7064
19.1934 | -106.6954
19.1972 | -106.6828
19.2009 | -106.6702
19.2046 | -106.6589
19.2084 | -106.648
19.2121 | -106.6352
19.2159 | -106.6228
19.2196 | -106.6101
19.2234 | -106.5959
19.2271 | -106.5861
19.2308 | -106.577
19.2346 | -106.5686
19.2383 | -106.5594
19.2421 | -106.55
19.2458 | -106.5432
19.2496 | -106.5377
19.2533 | -106.5305
19.257 | -106.5251
19.2608 | -106.5216
19.2645 | -106.5179
19.2683 | -106.5164
19.272 | -106.5157
19.2758 | -106.5162
19.2795 | -106.5201
19.2832 | -106.5275
19.287 | -106.5313
19.2907 | -106.5353
19.2945 | -106.5417
19.2982 | -106.5516
19.302 | -106.5591
19.3057 | -106.5683
19.3094 | -106.5781
19.3132 | -106.5906
19.3169 | -106.6018
19.3207 | -106.6102
19.3244 | -106.6179
19.3282 | -106.6258
19.3319 | -106.6334
19.3356 | -106.6408
19.3394 | -106.6507
19.3431 | -106.6596
19.3469 | -106.6679
19.3506 | -106.6785
19.3544 | -106.6882
19.3581 | -106.6977
19.3618 | -106.7058
19.3656 | -106.7132
19.3693 | -106.7199
19.3731 | -106.7244
19.3768 | -106.7296
19.3806 | -106.7338
19.3843 | -106.739
19.388 | -106.7454
19.3918 | -106.7526
19.3955 | -106.7552
19.3993 | -106.7571
19.403 | -106.7613
19.4068 | -106.7664
19.4105 | -106.7689
19.4142 | -106.7702
19.418 | -106.7716
19.4217 | -106.7731
19.4255 | -106.7762
19.4292 | -106.7778
19.433 | -106.7797
19.4367 | -106.7794
19.4404 | -106.779
19.4442 | -106.7759
19.4479 | -106.7738
19.4517 | -106.7727
19.4554 | -106.7719
19.4592 | -106.771
19.4629 | -106.7674
19.4666 | -106.7626
19.4704 | -106.7565
19.4741 | -106.752
19.4779 | -106.7482
19.4816 | -106.7406
19.4854 | -106.7333
19.4891 | -106.726
19.4928 | -106.7187
19.4966 | -106.7113
19.5003 | -106.7044
19.5041 | -106.6955
19.5078 | -106.6858
19.5116 | -106.6771
19.5153 | -106.6678
19.519 | -106.661
19.5228 | -106.6542
19.5265 | -106.6456
19.5303 | -106.6389
19.534 | -106.6301
19.5378 | -106.6198
19.5415 | -106.6097
19.5452 | -106.6017
19.549 | -106.5921
19.5527 | -106.5829
19.5565 | -106.5737
19.5602 | -106.5644
19.564 | -106.5604
19.5677 | -106.557
19.5714 | -106.5533
19.5752 | -106.5491
19.5789 | -106.5468
19.5827 | -106.5466
19.5864 | -106.5478
19.5902 | -106.5474
19.5939 | -106.5477
19.5976 | -106.5524
19.6014 | -106.5566
19.6051 | -106.5606
19.6089 | -106.5655
19.6126 | -106.5697
19.6164 | -106.5746
19.6201 | -106.5822
19.6238 | -106.5881
19.6276 | -106.5923
19.6313 | -106.5986
19.6351 | -106.6068
19.6388 | -106.6158
19.6426 | -106.6226
19.6463 | -106.6299
19.65 | -106.6379
19.6538 | -106.6449
19.6575 | -106.6519
19.6613 | -106.6592
19.665 | -106.6666
19.6688 | -106.6727
19.6725 | -106.6799
19.6762 | -106.6871
19.68 | -106.6936
19.6837 | -106.6988
19.6875 | -106.703
19.6912 | -106.7084
19.695 | -106.7144
19.6987 | -106.7213
19.7024 | -106.7278
19.7062 | -106.7339
19.7099 | -106.7384
19.7137 | -106.745
19.7174 | -106.7506
19.7212 | -106.756
19.7249 | -106.7628
19.7286 | -106.7679
19.7324 | -106.7737
19.7361 | -106.7802
19.7399 | -106.7843
19.7436 | -106.7921
19.7474 | -106.8002
19.7511 | -106.807
19.7548 | -106.8132
19.7586 | -106.8194
19.7623 | -106.8264
19.7661 | -106.8312
19.7698 | -106.839
19.7736 | -106.8456
19.7773 | -106.8504
19.781 | -106.8566
19.7848 | -106.862
19.7885 | -106.8675
19.7923 | -106.8733
19.796 | -106.8793
19.7998 | -106.8837
19.8035 | -106.8849
19.8072 | -106.8841
19.811 | -106.8854
19.8147 | -106.8861
19.8185 | -106.886
19.8222 | -106.8861
19.826 | -106.8833
19.8297 | -106.8802
19.8334 | -106.8758
19.8372 | -106.8702
19.8409 | -106.8642
19.8447 | -106.8577
19.8484 | -106.8497
19.8522 | -106.8397
19.8559 | -106.8301
19.8596 | -106.8197
19.8634 | -106.8078
19.8671 | -106.7962
19.8709 | -106.7861
19.8746 | -106.7788
19.8784 | -106.7688
19.8821 | -106.7588
19.8858 | -106.7484
19.8896 | -106.7355
19.8933 | -106.725
19.8971 | -106.7156
19.9008 | -106.706
19.9046 | -106.6992
19.9083 | -106.6936
19.912 | -106.6882
19.9158 | -106.6804
19.9195 | -106.6726
19.9233 | -106.6682
19.927 | -106.6648
19.9308 | -106.6613
19.9345 | -106.6586
19.9382 | -106.6558
19.942 | -106.6529
19.9457 | -106.6508
19.9495 | -106.6505
19.9532 | -106.6509
19.957 | -106.6498
19.9607 | -106.6483
19.9644 | -106.6477
19.9682 | -106.6489
19.9719 | -106.6496
19.9757 | -106.6508
19.9794 | -106.6482
19.9832 | -106.647
19.9869 | -106.645
19.9906 | -106.6431
19.9944 | -106.6424
19.9981 | -106.6419
20.0019 | -106.6408
20.0056 | -106.6398
20.0094 | -106.6375
20.0131 | -106.6363
20.0168 | -106.6346
20.0206 | -106.6329
20.0243 | -106.6312
20.0281 | -106.6305
20.0318 | -106.6315
20.0356 | -106.6312
20.0393 | -106.6293
20.043 | -106.6276
20.0468 | -106.6259
20.0505 | -106.6255
20.0543 | -106.6242
20.058 | -106.6243
20.0618 | -106.6244
20.0655 | -106.6223
20.0692 | -106.6213
20.073 | -106.6199
20.0767 | -106.6177
20.0805 | -106.6164
20.0842 | -106.6161
20.088 | -106.617
20.0917 | -106.6194
20.0954 | -106.618
20.0992 | -106.6174
20.1029 | -106.6187
20.1067 | -106.6186
20.1104 | -106.6169
20.1142 | -106.6171
20.1179 | -106.6187
20.1216 | -106.6207
20.1254 | -106.6235
20.1291 | -106.6262
20.1329 | -106.6292
20.1366 | -106.6331
20.1404 | -106.6363
20.1441 | -106.6398
20.1478 | -106.6441
20.1516 | -106.6493
20.1553 | -106.6528
20.1591 | -106.6548
20.1628 | -106.66
20.1666 | -106.6645
20.1703 | -106.6684
20.174 | -106.6734
20.1778 | -106.6765
20.1815 | -106.6782
20.1853 | -106.6804
20.189 | -106.6833
20.1928 | -106.6838
20.1965 | -106.6819
20.2002 | -106.6819
20.204 | -106.6823
20.2077 | -106.6819
20.2115 | -106.6803
20.2152 | -106.6811
20.219 | -106.6821
20.2227 | -106.6819
20.2264 | -106.6825
20.2302 | -106.6822
20.2339 | -106.6814
20.2377 | -106.6793
20.2414 | -106.6799
20.2452 | -106.6805
20.2489 | -106.6796
20.2526 | -106.6788
20.2564 | -106.6768
20.2601 | -106.6727
20.2639 | -106.6691
20.2676 | -106.6687
20.2714 | -106.667
20.2751 | -106.6648
20.2788 | -106.6622
20.2826 | -106.6594
20.2863 | -106.6547
20.2901 | -106.6516
20.2938 | -106.648
20.2976 | -106.6438
20.3013 | -106.6418
20.305 | -106.6382
20.3088 | -106.6341
20.3125 | -106.6313
20.3163 | -106.6279
20.32 | -106.6244
20.3238 | -106.6228
20.3275 | -106.619
20.3312 | -106.6144
20.335 | -106.6099
20.3387 | -106.6053
20.3425 | -106.6028
20.3462 | -106.5997
20.35 | -106.5971
20.3537 | -106.5957
20.3574 | -106.5945
20.3612 | -106.5939
20.3649 | -106.5936
20.3687 | -106.5944
20.3724 | -106.5932
20.3762 | -106.593
20.3799 | -106.5909
20.3836 | -106.5893
20.3874 | -106.5878
20.3911 | -106.5864
20.3949 | -106.5871
20.3986 | -106.5879
20.4024 | -106.5849
20.4061 | -106.5824
20.4098 | -106.5807
20.4136 | -106.5792
20.4173 | -106.577
20.4211 | -106.5761
20.4248 | -106.5748
20.4286 | -106.5736
20.4323 | -106.5715
20.436 | -106.5678
20.4398 | -106.5625
20.4435 | -106.559
20.4473 | -106.5557
20.451 | -106.5511
20.4548 | -106.5466
20.4585 | -106.5434
20.4622 | -106.5396
20.466 | -106.537
20.4697 | -106.5345
20.4735 | -106.5307
20.4772 | -106.525
20.481 | -106.5192
20.4847 | -106.513
20.4884 | -106.5079
20.4922 | -106.5012
20.4959 | -106.4938
20.4997 | -106.4847
20.5034 | -106.4766
20.5072 | -106.4691
20.5109 | -106.4627
20.5146 | -106.4568
20.5184 | -106.4512
20.5221 | -106.4477
20.5259 | -106.4446
20.5296 | -106.4417
20.5334 | -106.437
20.5371 | -106.4318
20.5408 | -106.4291
20.5446 | -106.4271
20.5483 | -106.4254
20.5521 | -106.426
20.5558 | -106.4279
20.5596 | -106.4272
20.5633 | -106.4266
20.567 | -106.4265
20.5708 | -106.4276
20.5745 | -106.4295
20.5783 | -106.4333
20.582 | -106.4374
20.5858 | -106.4432
20.5895 | -106.4484
20.5932 | -106.451
20.597 | -106.4536
20.6007 | -106.4572
20.6045 | -106.4613
20.6082 | -106.4639
20.612 | -106.4667
20.6157 | -106.4687
20.6194 | -106.4732
20.6232 | -106.4774
20.6269 | -106.481
20.6307 | -106.485
20.6344 | -106.4868
20.6382 | -106.4875
20.6419 | -106.4889
20.6456 | -106.4899
20.6494 | -106.4911
20.6531 | -106.4937
20.6569 | -106.4949
20.6606 | -106.4942
20.6644 | -106.4942
20.6681 | -106.4937
20.6718 | -106.4933
20.6756 | -106.4953
20.6793 | -106.498
20.6831 | -106.4968
20.6868 | -106.4914
20.6906 | -106.4899
20.6943 | -106.492
20.698 | -106.4908
20.7018 | -106.4909
20.7055 | -106.4907
20.7093 | -106.4903
20.713 | -106.4869
20.7168 | -106.4834
20.7205 | -106.4795
20.7242 | -106.476
20.728 | -106.4729
20.7317 | -106.4714
20.7355 | -106.4707
20.7392 | -106.4689
20.743 | -106.4664
20.7467 | -106.4634
20.7504 | -106.4602
20.7542 | -106.4585
20.7579 | -106.4559
20.7617 | -106.4521
20.7654 | -106.4468
20.7692 | -106.4432
20.7729 | -106.4411
20.7766 | -106.4367
20.7804 | -106.43
20.7841 | -106.4252
20.7879 | -106.4204
20.7916 | -106.4144
20.7954 | -106.4061
20.7991 | -106.3976
20.8028 | -106.391
20.8066 | -106.3835
20.8103 | -106.3791
20.8141 | -106.3719
20.8178 | -106.361
20.8216 | -106.3523
20.8253 | -106.3445
20.829 | -106.3358
20.8328 | -106.329
20.8365 | -106.3215
20.8403 | -106.3141
20.844 | -106.3071
20.8478 | -106.3001
20.8515 | -106.2931
20.8552 | -106.2884
20.859 | -106.2843
20.8627 | -106.2815
20.8665 | -106.2779
20.8702 | -106.2727
20.874 | -106.2687
20.8777 | -106.264
20.8814 | -106.2593
20.8852 | -106.256
20.8889 | -106.2529
20.8927 | -106.2505
20.8964 | -106.2479
20.9002 | -106.2447
20.9039 | -106.2427
20.9076 | -106.242
20.9114 | -106.2443
20.9151 | -106.2458
20.9189 | -106.2468
20.9226 | -106.2497
20.9264 | -106.2512
20.9301 | -106.2526
20.9338 | -106.2578
20.9376 | -106.2645
20.9413 | -106.2713
20.9451 | -106.2776
20.9488 | -106.2828
20.9526 | -106.2909
20.9563 | -106.3009
20.96 | -106.3109
20.9638 | -106.3221
20.9675 | -106.3351
20.9713 | -106.3469
20.975 | -106.3569
20.9788 | -106.3678
20.9825 | -106.3812
20.9862 | -106.392
20.99 | -106.4026
20.9937 | -106.4166
20.9975 | -106.4284
21.0012 | -106.4397
21.005 | -106.454
21.0087 | -106.4683
21.0124 | -106.4834
21.0162 | -106.4983
21.0199 | -106.5122
21.0237 | -106.5266
21.0274 | -106.5391
21.0312 | -106.5498
21.0349 | -106.5617
21.0386 | -106.5735
21.0424 | -106.5857
21.0461 | -106.596
21.0499 | -106.6065
21.0536 | -106.6175
21.0574 | -106.6267
21.0611 | -106.6368
21.0648 | -106.6483
21.0686 | -106.6595
21.0723 | -106.6702
21.0761 | -106.6818
21.0798 | -106.6931
21.0836 | -106.7027
21.0873 | -106.7114
21.091 | -106.7205
21.0948 | -106.7282
21.0985 | -106.7352
21.1023 | -106.7376
21.106 | -106.7393
21.1098 | -106.7416
21.1135 | -106.7467
21.1172 | -106.7515
21.121 | -106.7531
21.1247 | -106.7532
21.1285 | -106.7538
21.1322 | -106.7523
21.136 | -106.7489
21.1397 | -106.7454
21.1434 | -106.7393
21.1472 | -106.731
21.1509 | -106.7237
21.1547 | -106.7144
21.1584 | -106.7033
21.1622 | -106.6912
21.1659 | -106.6804
21.1696 | -106.6684
21.1734 | -106.6557
21.1771 | -106.6443
21.1809 | -106.6326
21.1846 | -106.6216
21.1884 | -106.611
21.1921 | -106.5984
21.1958 | -106.5835
21.1996 | -106.5704
21.2033 | -106.5579
21.2071 | -106.5449
21.2108 | -106.5303
21.2146 | -106.5159
21.2183 | -106.5013
21.222 | -106.4855
21.2258 | -106.4705
21.2295 | -106.454
21.2333 | -106.4363
21.237 | -106.4203
21.2408 | -106.4043
21.2445 | -106.3889
21.2482 | -106.3742
21.252 | -106.3609
21.2557 | -106.3465
21.2595 | -106.3324
21.2632 | -106.3172
21.267 | -106.3022
21.2707 | -106.2876
21.2744 | -106.2737
21.2782 | -106.2615
21.2819 | -106.2507
21.2857 | -106.2396
21.2894 | -106.227
21.2932 | -106.216
21.2969 | -106.2064
21.3006 | -106.1982
21.3044 | -106.1864
21.3081 | -106.1746
21.3119 | -106.1635
21.3156 | -106.1532
21.3194 | -106.1438
21.3231 | -106.1348
21.3268 | -106.1251
21.3306 | -106.1147
21.3343 | -106.1052
21.3381 | -106.0972
21.3418 | -106.0902
21.3456 | -106.0826
21.3493 | -106.0729
21.353 | -106.0648
21.3568 | -106.0581
21.3605 | -106.0497
21.3643 | -106.0427
21.368 | -106.0367
21.3718 | -106.0306
21.3755 | -106.0252
21.3792 | -106.0218
21.383 | -106.017
21.3867 | -106.0128
21.3905 | -106.0105
21.3942 | -106.0078
21.398 | -106.0055
21.4017 | -106.0029
21.4054 | -106.0039
21.4092 | -106.006
21.4129 | -106.0083
21.4167 | -106.0108
21.4204 | -106.0142
21.4242 | -106.0178
21.4279 | -106.0241
21.4316 | -106.0312
21.4354 | -106.038
21.4391 | -106.0466
21.4429 | -106.0533
21.4466 | -106.0644
21.4504 | -106.077
21.4541 | -106.0864
21.4578 | -106.0969
21.4616 | -106.1094
21.4653 | -106.1235
21.4691 | -106.1369
21.4728 | -106.1509
21.4766 | -106.1643
21.4803 | -106.1788
21.484 | -106.1942
21.4878 | -106.2104
21.4915 | -106.2249
21.4953 | -106.2392
21.499 | -106.2552
21.5028 | -106.2707
21.5065 | -106.2867
21.5102 | -106.3016
21.514 | -106.3153
21.5177 | -106.3305
21.5215 | -106.3452
21.5252 | -106.3603
21.529 | -106.3747
21.5327 | -106.3895
21.5364 | -106.403
21.5402 | -106.4138
21.5439 | -106.4237
21.5477 | -106.4344
21.5514 | -106.4452
21.5552 | -106.452
21.5589 | -106.4581
21.5626 | -106.4655
21.5664 | -106.4713
21.5701 | -106.4764
21.5739 | -106.4801
21.5776 | -106.4843
21.5814 | -106.4872
21.5851 | -106.4882
21.5888 | -106.4887
21.5926 | -106.4895
21.5963 | -106.4837
21.6001 | -106.4817
21.6038 | -106.479
21.6076 | -106.4753
21.6113 | -106.4756
21.615 | -106.4781
21.6188 | -106.4783
21.6225 | -106.4774
21.6263 | -106.4727
21.63 | -106.4674
21.6338 | -106.4662
21.6375 | -106.4651
21.6412 | -106.4618
21.645 | -106.4567
21.6487 | -106.4503
21.6525 | -106.4461
21.6562 | -106.4435
21.66 | -106.4399
21.6637 | -106.4368
21.6674 | -106.4337
21.6712 | -106.4287
21.6749 | -106.4218
21.6787 | -106.4148
21.6824 | -106.4084
21.6862 | -106.4015
21.6899 | -106.3936
21.6936 | -106.3854
21.6974 | -106.3765
21.7011 | -106.3704
21.7049 | -106.3644
21.7086 | -106.3586
21.7124 | -106.3508
21.7161 | -106.3416
21.7198 | -106.3313
21.7236 | -106.3241
21.7273 | -106.3166
21.7311 | -106.3099
21.7348 | -106.3035
21.7386 | -106.2979
21.7423 | -106.293
21.746 | -106.2849
21.7498 | -106.2792
21.7535 | -106.2707
21.7573 | -106.2629
21.761 | -106.2576
21.7648 | -106.2499
21.7685 | -106.2408
21.7722 | -106.2329
21.776 | -106.2253
21.7797 | -106.2172
21.7835 | -106.2096
21.7872 | -106.2001
21.791 | -106.1909
21.7947 | -106.1821
21.7984 | -106.1752
21.8022 | -106.1705
21.8059 | -106.164
21.8097 | -106.158
21.8134 | -106.1499
21.8172 | -106.1418
21.8209 | -106.1341
21.8246 | -106.1268
21.8284 | -106.1208
21.8321 | -106.1125
21.8359 | -106.1037
21.8396 | -106.097
21.8434 | -106.0897
21.8471 | -106.0819
21.8508 | -106.0756
21.8546 | -106.0691
21.8583 | -106.0604
21.8621 | -106.0539
21.8658 | -106.0489
21.8696 | -106.0427
21.8733 | -106.0371
21.877 | -106.0317
21.8808 | -106.0276
21.8845 | -106.0224
21.8883 | -106.016
21.892 | -106.0094
21.8958 | -106.0049
21.8995 | -105.9997
21.9032 | -105.9959
21.907 | -105.9905
21.9107 | -105.9863
21.9145 | -105.9815
21.9182 | -105.977
21.922 | -105.9741
21.9257 | -105.9689
21.9294 | -105.9648
21.9332 | -105.961
21.9369 | -105.9601
21.9407 | -105.9594
21.9444 | -105.9553
21.9482 | -105.9521
21.9519 | -105.9516
21.9556 | -105.9527
21.9594 | -105.9538
21.9631 | -105.9539
21.9669 | -105.9533
21.9706 | -105.952
21.9744 | -105.9519
21.9781 | -105.9554
21.9818 | -105.9595
21.9856 | -105.9643
21.9893 | -105.9682
21.9931 | -105.9715
21.9968 | -105.9754
22.0006 | -105.9808
22.0043 | -105.9858
22.008 | -105.9915
22.0118 | -105.9949
22.0155 | -105.9989
22.0193 | -106.0031
22.023 | -106.0087
22.0268 | -106.015
22.0305 | -106.022
22.0342 | -106.0266
22.038 | -106.0324
22.0417 | -106.0379
22.0455 | -106.0419
22.0492 | -106.0462
22.053 | -106.049
22.0567 | -106.0536
22.0604 | -106.057
22.0642 | -106.0615
22.0679 | -106.0642
22.0717 | -106.0643
22.0754 | -106.0656
22.0792 | -106.0673
22.0829 | -106.0672
22.0866 | -106.0663
22.0904 | -106.0669
22.0941 | -106.0671
22.0979 | -106.0651
22.1016 | -106.0619
22.1054 | -106.0574
22.1091 | -106.0524
22.1128 | -106.0453
22.1166 | -106.0398
22.1203 | -106.0341
22.1241 | -106.0294
22.1278 | -106.0259
22.1316 | -106.022
22.1353 | -106.0142
22.139 | -106.0053
22.1428 | -105.9975
22.1465 | -105.9916
22.1503 | -105.9872
22.154 | -105.9824
22.1578 | -105.9768
22.1615 | -105.9693
22.1652 | -105.9617
22.169 | -105.9557
22.1727 | -105.9513
22.1765 | -105.9452
22.1802 | -105.9371
22.184 | -105.93
22.1877 | -105.9229
22.1914 | -105.9173
22.1952 | -105.9112
22.1989 | -105.9051
22.2027 | -105.9019
22.2064 | -105.8974
22.2102 | -105.8912
22.2139 | -105.8861
22.2176 | -105.8829
22.2214 | -105.8802
22.2251 | -105.8773
22.2289 | -105.8752
22.2326 | -105.874
22.2364 | -105.8695
22.2401 | -105.8671
22.2438 | -105.8648
22.2476 | -105.8619
22.2513 | -105.8585
22.2551 | -105.8556
22.2588 | -105.8545
22.2626 | -105.8548
22.2663 | -105.8547
22.27 | -105.8555
22.2738 | -105.8579
22.2775 | -105.8595
22.2813 | -105.8604
22.285 | -105.8624
22.2888 | -105.8639
22.2925 | -105.865
22.2962 | -105.8668
22.3 | -105.8708
22.3037 | -105.8758
22.3075 | -105.882
22.3112 | -105.8885
22.315 | -105.8954
22.3187 | -105.9037
22.3224 | -105.9111
22.3262 | -105.9197
22.3299 | -105.9287
22.3337 | -105.9373
22.3374 | -105.9454
22.3412 | -105.9542
22.3449 | -105.9615
22.3486 | -105.969
22.3524 | -105.9757
22.3561 | -105.9832
22.3599 | -105.9932
22.3636 | -106.0028
22.3674 | -106.0153
22.3711 | -106.0273
22.3748 | -106.0388
22.3786 | -106.0489
22.3823 | -106.0592
22.3861 | -106.0701
22.3898 | -106.0816
22.3936 | -106.0906
22.3973 | -106.0996
22.401 | -106.1077
22.4048 | -106.1148
22.4085 | -106.1234
22.4123 | -106.1342
22.416 | -106.1432
22.4198 | -106.1528
22.4235 | -106.1617
22.4272 | -106.1703
22.431 | -106.1778
22.4347 | -106.185
22.4385 | -106.1943
22.4422 | -106.2024
22.446 | -106.21
22.4497 | -106.2196
22.4534 | -106.2299
22.4572 | -106.2365
22.4609 | -106.2437
22.4647 | -106.2499
22.4684 | -106.256
22.4722 | -106.2612
22.4759 | -106.2672
22.4796 | -106.2737
22.4834 | -106.2802
22.4871 | -106.2867
22.4909 | -106.2917
22.4946 | -106.2985
22.4984 | -106.3065
22.5021 | -106.3135
22.5058 | -106.3182
22.5096 | -106.3182
22.5133 | -106.3197
22.5171 | -106.3223
22.5208 | -106.3235
22.5246 | -106.3246
22.5283 | -106.3259
22.532 | -106.3252
22.5358 | -106.3201
22.5395 | -106.3163
22.5433 | -106.3111
22.547 | -106.3026
22.5508 | -106.2937
22.5545 | -106.286
22.5582 | -106.2777
22.562 | -106.2695
22.5657 | -106.2607
22.5695 | -106.2489
22.5732 | -106.2325
22.577 | -106.2172
22.5807 | -106.1998
22.5844 | -106.1794
22.5882 | -106.1561
22.5919 | -106.1328
22.5957 | -106.1124
22.5994 | -106.0905
22.6032 | -106.0668
22.6069 | -106.0419
22.6106 | -106.0166
22.6144 | -105.9907
22.6181 | -105.9624
22.6219 | -105.9341
22.6256 | -105.9053
22.6294 | -105.8754
22.6331 | -105.8439
22.6368 | -105.8132
22.6406 | -105.7824
22.6443 | -105.751
22.6481 | -105.7199
22.6518 | -105.6893
22.6556 | -105.6601
22.6593 | -105.6318
22.663 | -105.6056
22.6668 | -105.5782
22.6705 | -105.552
22.6743 | -105.5279
22.678 | -105.5092
22.6818 | -105.4844
22.6855 | -105.4619
22.6892 | -105.4403
22.693 | -105.4205
22.6967 | -105.4005
22.7005 | -105.3821
22.7042 | -105.3667
22.708 | -105.3552
22.7117 | -105.3431
22.7154 | -105.3335
22.7192 | -105.3247
22.7229 | -105.3172
22.7267 | -105.3116
22.7304 | -105.3061
22.7342 | -105.3016
22.7379 | -105.2991
22.7416 | -105.2982
22.7454 | -105.2988
22.7491 | -105.3025
22.7529 | -105.3072
22.7566 | -105.3126
22.7604 | -105.3181
22.7641 | -105.3257
22.7678 | -105.3363
22.7716 | -105.3467
22.7753 | -105.3568
22.7791 | -105.3677
22.7828 | -105.3806
22.7866 | -105.3935
22.7903 | -105.4053
22.794 | -105.4191
22.7978 | -105.4329
22.8015 | -105.4472
22.8053 | -105.4616
22.809 | -105.4761
22.8128 | -105.4918
22.8165 | -105.5064
22.8202 | -105.52
22.824 | -105.5337
22.8277 | -105.5484
22.8315 | -105.5651
22.8352 | -105.5804
22.839 | -105.5944
22.8427 | -105.6067
22.8464 | -105.6195
22.8502 | -105.6329
22.8539 | -105.6464
22.8577 | -105.6592
22.8614 | -105.6708
22.8652 | -105.6811
22.8689 | -105.694
22.8726 | -105.7039
22.8764 | -105.7121
22.8801 | -105.7193
22.8839 | -105.7255
22.8876 | -105.7313
22.8914 | -105.738
22.8951 | -105.7448
22.8988 | -105.7477
22.9026 | -105.7513
22.9063 | -105.7566
22.9101 | -105.7622
22.9138 | -105.7641
22.9176 | -105.7631
22.9213 | -105.7653
22.925 | -105.7666
22.9288 | -105.7684
22.9325 | -105.77
22.9363 | -105.7699
22.94 | -105.769
22.9438 | -105.7687
22.9475 | -105.7686
22.9512 | -105.7697
22.955 | -105.7728
22.9587 | -105.7742
22.9625 | -105.7756
22.9662 | -105.7789
22.97 | -105.7802
22.9737 | -105.7803
22.9774 | -105.7814
22.9812 | -105.7837
22.9849 | -105.7859
22.9887 | -105.7902
22.9924 | -105.7947
22.9962 | -105.7967
22.9999 | -105.8
23.0036 | -105.8062
23.0074 | -105.8118
23.0111 | -105.816
23.0149 | -105.8212
23.0186 | -105.8252
23.0224 | -105.8296
23.0261 | -105.8354
23.0298 | -105.8415
23.0336 | -105.8467
23.0373 | -105.852
23.0411 | -105.858
23.0448 | -105.8629
23.0486 | -105.8699
23.0523 | -105.876
23.056 | -105.8805
23.0598 | -105.8849
23.0635 | -105.8902
23.0673 | -105.8939
23.071 | -105.8969
23.0748 | -105.8977
23.0785 | -105.8977
23.0822 | -105.8971
23.086 | -105.8954
23.0897 | -105.8933
23.0935 | -105.8881
23.0972 | -105.8851
23.101 | -105.8827
23.1047 | -105.8787
23.1085 | -105.8712
23.1122 | -105.8653
23.1159 | -105.8576
23.1197 | -105.847
23.1234 | -105.8388
23.1272 | -105.8315
23.1309 | -105.8209
23.1347 | -105.8079
23.1384 | -105.7958
23.1421 | -105.7847
23.1459 | -105.7731
23.1496 | -105.7617
23.1534 | -105.75
23.1571 | -105.7386
23.1608 | -105.7273
23.1646 | -105.7166
23.1683 | -105.7085
23.1721 | -105.6992
23.1758 | -105.6882
23.1796 | -105.6799
23.1833 | -105.6727
23.187 | -105.6648
23.1908 | -105.6588
23.1945 | -105.6536
23.1983 | -105.6501
23.202 | -105.6489
23.2058 | -105.6478
23.2095 | -105.6461
23.2132 | -105.6474
23.217 | -105.6563
23.2207 | -105.6576
23.2245 | -105.6619
23.2282 | -105.6658
23.232 | -105.6706
23.2357 | -105.6742
23.2395 | -105.678
23.2432 | -105.6858
23.2469 | -105.6949
23.2507 | -105.7018
23.2544 | -105.7075
23.2582 | -105.714
23.2619 | -105.7201
23.2657 | -105.7272
23.2694 | -105.7357
23.2731 | -105.7444
23.2769 | -105.753
23.2806 | -105.7611
23.2844 | -105.7681
23.2881 | -105.7765
23.2919 | -105.7849
23.2956 | -105.7926
23.2993 | -105.8004
23.3031 | -105.8092
23.3068 | -105.8159
23.3106 | -105.8206
23.3143 | -105.8242
23.318 | -105.8282
23.3218 | -105.8325
23.3255 | -105.836
23.3293 | -105.8389
23.333 | -105.8402
23.3368 | -105.8397
23.3405 | -105.839
23.3442 | -105.8381
23.348 | -105.8371
23.3517 | -105.8359
23.3555 | -105.8327
23.3592 | -105.8266
23.363 | -105.819
23.3667 | -105.8129
23.3705 | -105.8066
23.3742 | -105.8002
23.3779 | -105.7927
23.3817 | -105.7849
23.3854 | -105.7788
23.3892 | -105.7728
23.3929 | -105.7658
23.3967 | -105.7584
23.4004 | -105.7496
23.4041 | -105.7418
23.4079 | -105.7359
23.4116 | -105.73
23.4154 | -105.7245
23.4191 | -105.7188
23.4229 | -105.7135
23.4266 | -105.7056
23.4303 | -105.6965
23.4341 | -105.6908
23.4378 | -105.6857
23.4416 | -105.6792
23.4453 | -105.675
23.449 | -105.672
23.4528 | -105.6658
23.4565 | -105.6614
23.4603 | -105.6581
23.464 | -105.6548
23.4678 | -105.6513
23.4715 | -105.6483
23.4752 | -105.6471
23.479 | -105.6434
23.4827 | -105.6407
23.4865 | -105.6393
23.4902 | -105.6368
23.494 | -105.6352
23.4977 | -105.6335
23.5014 | -105.6327
23.5052 | -105.6342
23.5089 | -105.6347
23.5127 | -105.6368
23.5164 | -105.6408
23.5202 | -105.644
23.5239 | -105.6461
23.5277 | -105.6491
23.5314 | -105.6511
23.5351 | -105.655
23.5389 | -105.6592
23.5426 | -105.6605
23.5464 | -105.6626
23.5501 | -105.6675
23.5539 | -105.6734
23.5576 | -105.6777
23.5613 | -105.6834
23.5651 | -105.6884
23.5688 | -105.6956
23.5726 | -105.7056
23.5763 | -105.7134
23.5801 | -105.7206
23.5838 | -105.7295
23.5875 | -105.7408
23.5913 | -105.7516
23.595 | -105.7616
23.5988 | -105.771
23.6025 | -105.781
23.6063 | -105.7904
23.61 | -105.7999
23.6137 | -105.8102
23.6175 | -105.8217
23.6212 | -105.8329
23.625 | -105.8443
23.6287 | -105.8525
23.6325 | -105.8615
23.6362 | -105.8705
23.6399 | -105.8796
23.6437 | -105.8884
23.6474 | -105.8984
23.6512 | -105.9074
23.6549 | -105.9156
23.6587 | -105.9265
23.6624 | -105.9359
23.6661 | -105.9448
23.6699 | -105.9543
23.6736 | -105.9632
23.6774 | -105.9713
23.6811 | -105.9813
23.6849 | -105.9903
23.6886 | -105.9977
23.6923 | -106.0055
23.6961 | -106.0117
23.6998 | -106.0179
23.7036 | -106.0235
23.7073 | -106.0298
23.7111 | -106.0346
23.7148 | -106.0384
23.7185 | -106.0427
23.7223 | -106.0464
23.726 | -106.0475
23.7298 | -106.0493
23.7335 | -106.053
23.7373 | -106.056
23.741 | -106.0565
23.7447 | -106.0544
23.7485 | -106.0545
23.7522 | -106.0577
23.756 | -106.0574
23.7597 | -106.06
23.7635 | -106.0627
23.7672 | -106.0658
23.7709 | -106.068
23.7747 | -106.0708
23.7784 | -106.0735
23.7822 | -106.0748
23.7859 | -106.0747
23.7897 | -106.0758
23.7934 | -106.0745
23.7971 | -106.0741
23.8009 | -106.0753
23.8046 | -106.0756
23.8084 | -106.0757
23.8121 | -106.0763
23.8159 | -106.0799
23.8196 | -106.0836
23.8233 | -106.0845
23.8271 | -106.0857
23.8308 | -106.0875
23.8346 | -106.0891
23.8383 | -106.0902
23.8421 | -106.0907
23.8458 | -106.0928
23.8495 | -106.0954
23.8533 | -106.0955
23.857 | -106.0965
23.8608 | -106.0992
23.8645 | -106.1004
23.8683 | -106.1021
23.872 | -106.1046
23.8757 | -106.1087
23.8795 | -106.1157
23.8832 | -106.1239
23.887 | -106.1308
23.8907 | -106.1346
23.8945 | -106.1385
23.8982 | -106.1441
23.9019 | -106.1494
23.9057 | -106.1544
23.9094 | -106.1569
23.9132 | -106.1598
23.9169 | -106.1651
23.9207 | -106.171
23.9244 | -106.1777
23.9281 | -106.1829
23.9319 | -106.1884
23.9356 | -106.1931
23.9394 | -106.1969
23.9431 | -106.2002
23.9469 | -106.203
23.9506 | -106.2038
23.9543 | -106.2059
23.9581 | -106.2084
23.9618 | -106.2101
23.9656 | -106.2131
23.9693 | -106.2154
23.9731 | -106.217
23.9768 | -106.216
23.9805 | -106.2157
23.9843 | -106.2137
23.988 | -106.2121
23.9918 | -106.2105
23.9955 | -106.2074
23.9993 | -106.2039
24.003 | -106.2019
24.0067 | -106.1992
24.0105 | -106.1971
24.0142 | -106.1954
24.018 | -106.1924
24.0217 | -106.1887
24.0255 | -106.1872
24.0292 | -106.1861
24.0329 | -106.1878
24.0367 | -106.1882
24.0404 | -106.1862
24.0442 | -106.1838
24.0479 | -106.1847
24.0517 | -106.1862
24.0554 | -106.1865
24.0591 | -106.1871
24.0629 | -106.1893
24.0666 | -106.1921
24.0704 | -106.1934
24.0741 | -106.196
24.0779 | -106.2011
24.0816 | -106.2076
24.0853 | -106.2098
24.0891 | -106.2101
24.0928 | -106.2141
24.0966 | -106.219
24.1003 | -106.2242
24.1041 | -106.2314
24.1078 | -106.2386
24.1115 | -106.2451
24.1153 | -106.2513
24.119 | -106.2586
24.1228 | -106.2659
24.1265 | -106.2741
24.1303 | -106.2835
24.134 | -106.2946
24.1377 | -106.3049
24.1415 | -106.3157
24.1452 | -106.3269
24.149 | -106.3376
24.1527 | -106.3478
24.1565 | -106.3569
24.1602 | -106.3667
24.1639 | -106.3781
24.1677 | -106.3908
24.1714 | -106.4013
24.1752 | -106.4126
24.1789 | -106.4254
24.1827 | -106.4366
24.1864 | -106.4468
24.1901 | -106.4572
24.1939 | -106.4695
24.1976 | -106.4807
24.2014 | -106.4918
24.2051 | -106.5052
24.2089 | -106.518
24.2126 | -106.5282
24.2163 | -106.5371
24.2201 | -106.5451
24.2238 | -106.5511
24.2276 | -106.558
24.2313 | -106.567
24.2351 | -106.5753
24.2388 | -106.5817
24.2425 | -106.588
24.2463 | -106.5936
24.25 | -106.5991
24.2538 | -106.6041
24.2575 | -106.6079
24.2613 | -106.6102
24.265 | -106.6118
24.2687 | -106.6143
24.2725 | -106.6168
24.2762 | -106.621
24.28 | -106.6251
24.2837 | -106.6278
24.2875 | -106.629
24.2912 | -106.6325
24.2949 | -106.6325
24.2987 | -106.632
24.3024 | -106.6301
24.3062 | -106.6276
24.3099 | -106.6278
24.3137 | -106.6288
24.3174 | -106.6302
24.3211 | -106.6322
24.3249 | -106.6335
24.3286 | -106.6352
24.3324 | -106.6368
24.3361 | -106.6381
24.3399 | -106.6382
24.3436 | -106.6386
24.3473 | -106.6377
24.3511 | -106.6373
24.3548 | -106.6363
24.3586 | -106.6343
24.3623 | -106.6322
24.3661 | -106.6319
24.3698 | -106.6287
24.3735 | -106.6252
24.3773 | -106.6225
24.381 | -106.6204
24.3848 | -106.6203
24.3885 | -106.6197
24.3923 | -106.6175
24.396 | -106.615
24.3997 | -106.6131
24.4035 | -106.6118
24.4072 | -106.612
24.411 | -106.6104
24.4147 | -106.6099
24.4185 | -106.6078
24.4222 | -106.6044
24.4259 | -106.6027
24.4297 | -106.6046
24.4334 | -106.6033
24.4372 | -106.5995
24.4409 | -106.5986
24.4447 | -106.5986
24.4484 | -106.5988
24.4521 | -106.5992
24.4559 | -106.5991
24.4596 | -106.6006
24.4634 | -106.6026
24.4671 | -106.6059
24.4709 | -106.6088
24.4746 | -106.6124
24.4783 | -106.6157
24.4821 | -106.6201
24.4858 | -106.6243
24.4896 | -106.6286
24.4933 | -106.6319
24.4971 | -106.6354
24.5008 | -106.6398
24.5045 | -106.6434
24.5083 | -106.6477
24.512 | -106.6511
24.5158 | -106.6562
24.5195 | -106.6612
24.5233 | -106.6647
24.527 | -106.6677
24.5307 | -106.6697
24.5345 | -106.6724
24.5382 | -106.6754
24.542 | -106.6792
24.5457 | -106.679
24.5495 | -106.6781
24.5532 | -106.6776
24.5569 | -106.6768
24.5607 | -106.6766
24.5644 | -106.6737
24.5682 | -106.6671
24.5719 | -106.6597
24.5757 | -106.6506
24.5794 | -106.6363
24.5831 | -106.5981
24.5869 | -106.5509
24.5906 | -106.5028
24.5944 | -106.4499
24.5981 | -106.3832
24.6019 | -106.3028
24.6056 | -106.2263
24.6093 | -106.1609
24.6131 | -106.0959
24.6168 | -106.0296
24.6206 | -105.9616
24.6243 | -105.8855
24.6281 | -105.8062
24.6318 | -105.7258
24.6355 | -105.641
24.6393 | -105.5544
24.643 | -105.4648
24.6468 | -105.3692
24.6505 | -105.2727
24.6543 | -105.1755
24.658 | -105.0742
24.6617 | -104.97
24.6655 | -104.8639
24.6692 | -104.7597
24.673 | -104.6535
24.6767 | -104.5466
24.6805 | -104.4408
24.6842 | -104.3351
24.6879 | -104.2318
24.6917 | -104.1306
24.6954 | -104.0317
24.6992 | -103.9352
24.7029 | -103.8414
24.7067 | -103.7499
24.7104 | -103.6622
24.7141 | -103.5781
24.7179 | -103.4985
24.7216 | -103.4227
24.7254 | -103.351
24.7291 | -103.2845
24.7329 | -103.2231
24.7366 | -103.1658
24.7403 | -103.1138
24.7441 | -103.0687
24.7478 | -103.0289
24.7516 | -102.9949
24.7553 | -102.9664
24.7591 | -102.9443
24.7628 | -102.9276
24.7665 | -102.9173
24.7703 | -102.9132
24.774 | -102.9153
24.7778 | -102.9223
24.7815 | -102.9361
24.7853 | -102.9562
24.789 | -102.9808
24.7927 | -103.0106
24.7965 | -103.0457
24.8002 | -103.0874
24.804 | -103.135
24.8077 | -103.1868
24.8115 | -103.2435
24.8152 | -103.3043
24.8189 | -103.3696
24.8227 | -103.4389
24.8264 | -103.5121
24.8302 | -103.5887
24.8339 | -103.6675
24.8377 | -103.7586
24.8414 | -103.8403
24.8451 | -103.9245
24.8489 | -104.0108
24.8526 | -104.0984
24.8564 | -104.1866
24.8601 | -104.2754
24.8639 | -104.3653
24.8676 | -104.4551
24.8713 | -104.5434
24.8751 | -104.632
24.8788 | -104.72
24.8826 | -104.8072
24.8863 | -104.8918
24.8901 | -104.9738
24.8938 | -105.0535
24.8975 | -105.1327
24.9013 | -105.2079
24.905 | -105.2793
24.9088 | -105.3467
24.9125 | -105.4102
24.9163 | -105.4724
24.92 | -105.5355
24.9237 | -105.5937
24.9275 | -105.6487
24.9312 | -105.7036
24.935 | -105.7701
24.9387 | -105.8342
24.9425 | -105.8886
24.9462 | -105.938
24.9499 | -105.9915
24.9537 | -106.0455
24.9574 | -106.0854
24.9612 | -106.1052
24.9649 | -106.1171
24.9687 | -106.1268
24.9724 | -106.1328
24.9761 | -106.1418
24.9799 | -106.1495
24.9836 | -106.1544
24.9874 | -106.1589
24.9911 | -106.1605
24.9949 | -106.161
24.9986 | -106.1654
25.0023 | -106.17
25.0061 | -106.1721
25.0098 | -106.1746
25.0136 | -106.1769
25.0173 | -106.1801
25.0211 | -106.1802
25.0248 | -106.1825
25.0285 | -106.1845
25.0323 | -106.1852
25.036 | -106.1872
25.0398 | -106.1894
25.0435 | -106.1926
25.0473 | -106.1931
25.051 | -106.1936
25.0547 | -106.1929
25.0585 | -106.1933
25.0622 | -106.1938
25.066 | -106.194
25.0697 | -106.1937
25.0735 | -106.1934
25.0772 | -106.1909
25.0809 | -106.1864
25.0847 | -106.1801
25.0884 | -106.1775
25.0922 | -106.1758
25.0959 | -106.1712
25.0997 | -106.1673
25.1034 | -106.1632
25.1071 | -106.1587
25.1109 | -106.1535
25.1146 | -106.1491
25.1184 | -106.1437
25.1221 | -106.1383
25.1259 | -106.1316
25.1296 | -106.1257
25.1333 | -106.1196
25.1371 | -106.1139
25.1408 | -106.1063
25.1446 | -106.0986
25.1483 | -106.0916
25.1521 | -106.0843
25.1558 | -106.0769
25.1595 | -106.0703
25.1633 | -106.0635
25.167 | -106.0578
25.1708 | -106.0526
25.1745 | -106.048
25.1783 | -106.0422
25.182 | -106.0369
25.1857 | -106.0338
25.1895 | -106.0297
25.1932 | -106.0254
25.197 | -106.0228
25.2007 | -106.0216
25.2045 | -106.0194
25.2082 | -106.0181
25.2119 | -106.018
25.2157 | -106.0178
25.2194 | -106.0187
25.2232 | -106.0201
25.2269 | -106.0222
25.2307 | -106.0254
25.2344 | -106.0289
25.2381 | -106.0312
25.2419 | -106.0326
25.2456 | -106.0362
25.2494 | -106.0416
25.2531 | -106.0453
25.2569 | -106.0513
25.2606 | -106.0591
25.2643 | -106.068
25.2681 | -106.0771
25.2718 | -106.0849
25.2756 | -106.0901
25.2793 | -106.0944
25.2831 | -106.0995
25.2868 | -106.1048
25.2905 | -106.1111
25.2943 | -106.1186
25.298 | -106.127
25.3018 | -106.1359
25.3055 | -106.1434
25.3093 | -106.1508
25.313 | -106.1576
25.3167 | -106.164
25.3205 | -106.1707
25.3242 | -106.1752
25.328 | -106.1796
25.3317 | -106.1822
25.3355 | -106.1858
25.3392 | -106.1891
25.3429 | -106.1906
25.3467 | -106.1923
25.3504 | -106.1951
25.3542 | -106.1963
25.3579 | -106.1957
25.3617 | -106.1954
25.3654 | -106.1945
25.3691 | -106.1946
25.3729 | -106.1921
25.3766 | -106.1831
25.3804 | -106.18
25.3841 | -106.1776
25.3879 | -106.1747
25.3916 | -106.1715
25.3953 | -106.1679
25.3991 | -106.1627
25.4028 | -106.157
25.4066 | -106.1531
25.4103 | -106.1487
25.4141 | -106.1449
25.4178 | -106.1405
25.4215 | -106.1344
25.4253 | -106.1278
25.429 | -106.1205
25.4328 | -106.1141
25.4365 | -106.1079
25.4403 | -106.1014
25.444 | -106.0969
25.4477 | -106.0909
25.4515 | -106.0835
25.4552 | -106.0761
25.459 | -106.0686
25.4627 | -106.0614
25.4665 | -106.0564
25.4702 | -106.0525
25.4739 | -106.0465
25.4777 | -106.0427
25.4814 | -106.0382
25.4852 | -106.0344
25.4889 | -106.0312
25.4927 | -106.0278
25.4964 | -106.025
25.5001 | -106.023
25.5039 | -106.0197
25.5076 | -106.0164
25.5114 | -106.0125
25.5151 | -106.0099
25.5189 | -106.0081
25.5226 | -106.0068
25.5263 | -106.0056
25.5301 | -106.0045
25.5338 | -106.0022
25.5376 | -106.0015
25.5413 | -106.0008
25.5451 | -106.0003
25.5488 | -106.0004
25.5525 | -106.0011
25.5563 | -106.0001
25.56 | -105.9991
25.5638 | -105.9993
25.5675 | -105.9979
25.5713 | -105.997
25.575 | -105.9974
25.5787 | -105.9959
25.5825 | -105.9955
25.5862 | -105.9977
25.59 | -105.9988
25.5937 | -105.9987
25.5975 | -106.0008
25.6012 | -106.0033
25.6049 | -106.0039
25.6087 | -106.0033
25.6124 | -106.0057
25.6162 | -106.0063
25.6199 | -106.006
25.6237 | -106.0071
25.6274 | -106.0085
25.6311 | -106.0093
25.6349 | -106.011
25.6386 | -106.0154
25.6424 | -106.019
25.6461 | -106.0227
25.6499 | -106.0255
25.6536 | -106.028
25.6573 | -106.0302
25.6611 | -106.0306
25.6648 | -106.0306
25.6686 | -106.0315
25.6723 | -106.0339
25.6761 | -106.0358
25.6798 | -106.0377
25.6835 | -106.0405
25.6873 | -106.0429
25.691 | -106.0451
25.6948 | -106.0462
25.6985 | -106.0466
25.7023 | -106.0486
25.706 | -106.0505
25.7097 | -106.0508
25.7135 | -106.0522
25.7172 | -106.0528
25.721 | -106.0533
25.7247 | -106.0547
25.7285 | -106.0539
25.7322 | -106.0541
25.7359 | -106.054
25.7397 | -106.0555
25.7434 | -106.0549
25.7472 | -106.052
25.7509 | -106.0508
25.7547 | -106.0493
25.7584 | -106.0477
25.7621 | -106.0457
25.7659 | -106.0431
25.7696 | -106.0393
25.7734 | -106.0359
25.7771 | -106.0318
25.7809 | -106.0286
25.7846 | -106.0266
25.7883 | -106.0231
25.7921 | -106.0189
25.7958 | -106.0166
25.7996 | -106.0131
25.8033 | -106.0105
25.8071 | -106.0068
25.8108 | -106.0029
25.8145 | -105.9976
25.8183 | -105.9919
25.822 | -105.9858
25.8258 | -105.9804
25.8295 | -105.9747
25.8333 | -105.9685
25.837 | -105.9635
25.8407 | -105.9585
25.8445 | -105.9546
25.8482 | -105.9499
25.852 | -105.9458
25.8557 | -105.9435
25.8595 | -105.9401
25.8632 | -105.9372
25.8669 | -105.9356
25.8707 | -105.9334
25.8744 | -105.9319
25.8782 | -105.9305
25.8819 | -105.9287
25.8857 | -105.9254
25.8894 | -105.9252
25.8931 | -105.9263
25.8969 | -105.9283
25.9006 | -105.9297
25.9044 | -105.9319
25.9081 | -105.9337
25.9119 | -105.9347
25.9156 | -105.9395
25.9193 | -105.9406
25.9231 | -105.9432
25.9268 | -105.9465
25.9306 | -105.9497
25.9343 | -105.9553
25.9381 | -105.961
25.9418 | -105.9658
25.9455 | -105.9714
25.9493 | -105.9767
25.953 | -105.9807
25.9568 | -105.9851
25.9605 | -105.9918
25.9643 | -105.9978
25.968 | -106.0049
25.9717 | -106.013
25.9755 | -106.0204
25.9792 | -106.0263
25.983 | -106.0325
25.9867 | -106.0379
25.9905 | -106.0444
25.9942 | -106.0516
25.9979 | -106.0601
26.0017 | -106.0677
26.0054 | -106.0742
26.0092 | -106.0804
26.0129 | -106.0868
26.0167 | -106.0941
26.0204 | -106.1011
26.0241 | -106.1076
26.0279 | -106.1146
26.0316 | -106.1217
26.0354 | -106.1283
26.0391 | -106.1348
26.0429 | -106.1407
26.0466 | -106.1454
26.0503 | -106.1506
26.0541 | -106.1548
26.0578 | -106.1591
26.0616 | -106.164
26.0653 | -106.169
26.0691 | -106.1743
26.0728 | -106.1794
26.0765 | -106.1826
26.0803 | -106.1843
26.084 | -106.188
26.0878 | -106.1908
26.0915 | -106.1935
26.0953 | -106.1978
26.099 | -106.2019
26.1027 | -106.2055
26.1065 | -106.2076
26.1102 | -106.2104
26.114 | -106.2131
26.1177 | -106.2139
26.1215 | -106.2147
26.1252 | -106.2166
26.1289 | -106.2199
26.1327 | -106.2225
26.1364 | -106.2264
26.1402 | -106.2284
26.1439 | -106.2299
26.1477 | -106.2302
26.1514 | -106.2301
26.1551 | -106.2319
26.1589 | -106.2326
26.1626 | -106.2292
26.1664 | -106.2298
26.1701 | -106.2303
26.1739 | -106.2293
26.1776 | -106.2258
26.1813 | -106.2252
26.1851 | -106.226
26.1888 | -106.2257
26.1926 | -106.2243
26.1963 | -106.2226
26.2001 | -106.2218
26.2038 | -106.2222
26.2075 | -106.221
26.2113 | -106.2181
26.215 | -106.2171
26.2188 | -106.2162
26.2225 | -106.2137
26.2263 | -106.2129
26.23 | -106.2132
26.2337 | -106.2134
26.2375 | -106.2135
26.2412 | -106.2122
26.245 | -106.2119
26.2487 | -106.2124
26.2525 | -106.2141
26.2562 | -106.2139
26.2599 | -106.212
26.2637 | -106.2103
26.2674 | -106.2101
26.2712 | -106.2119
26.2749 | -106.2132
26.2787 | -106.2136
26.2824 | -106.2149
26.2861 | -106.215
26.2899 | -106.2161
26.2936 | -106.2197
26.2974 | -106.223
26.3011 | -106.2258
26.3049 | -106.2259
26.3086 | -106.2255
26.3123 | -106.2256
26.3161 | -106.2266
26.3198 | -106.2273
26.3236 | -106.2281
26.3273 | -106.2301
26.3311 | -106.2324
26.3348 | -106.2318
26.3385 | -106.2311
26.3423 | -106.2316
26.346 | -106.2337
26.3498 | -106.2366
26.3535 | -106.2384
26.3573 | -106.24
26.361 | -106.24
26.3647 | -106.2396
26.3685 | -106.2386
26.3722 | -106.2395
26.376 | -106.2411
26.3797 | -106.2426
26.3835 | -106.2424
26.3872 | -106.2411
26.3909 | -106.241
26.3947 | -106.2413
26.3984 | -106.2422
26.4022 | -106.242
26.4059 | -106.2404
26.4097 | -106.2382
26.4134 | -106.2378
26.4171 | -106.2386
26.4209 | -106.2367
26.4246 | -106.2363
26.4284 | -106.2358
26.4321 | -106.233
26.4359 | -106.2314
26.4396 | -106.2291
26.4433 | -106.2258
26.4471 | -106.224
26.4508 | -106.222
26.4546 | -106.2193
26.4583 | -106.2169
26.4621 | -106.2143
26.4658 | -106.2108
26.4695 | -106.2079
26.4733 | -106.2043
26.477 | -106.2011
26.4808 | -106.1982
26.4845 | -106.1954
26.4883 | -106.1933
26.492 | -106.1925
26.4957 | -106.1906
26.4995 | -106.187
26.5032 | -106.1835
26.507 | -106.1793
26.5107 | -106.1736
26.5145 | -106.1705
26.5182 | -106.1686
26.5219 | -106.1678
26.5257 | -106.1642
26.5294 | -106.1597
26.5332 | -106.1573
26.5369 | -106.1567
26.5407 | -106.1538
26.5444 | -106.1506
26.5481 | -106.1495
26.5519 | -106.1496
26.5556 | -106.1473
26.5594 | -106.1449
26.5631 | -106.1448
26.5669 | -106.1446
26.5706 | -106.1433
26.5743 | -106.1401
26.5781 | -106.1372
26.5818 | -106.1355
26.5856 | -106.1344
26.5893 | -106.133
26.5931 | -106.1322
26.5968 | -106.1339
26.6005 | -106.1348
26.6043 | -106.1354
26.608 | -106.1365
26.6118 | -106.1378
26.6155 | -106.1391
26.6193 | -106.139
26.623 | -106.1403
26.6267 | -106.143
26.6305 | -106.1462
26.6342 | -106.1507
26.638 | -106.155
26.6417 | -106.1593
26.6455 | -106.1619
26.6492 | -106.1646
26.6529 | -106.1702
26.6567 | -106.1748
26.6604 | -106.1789
26.6642 | -106.1825
26.6679 | -106.1857
26.6717 | -106.189
26.6754 | -106.1926
26.6791 | -106.1983
26.6829 | -106.2035
26.6866 | -106.2093
26.6904 | -106.2134
26.6941 | -106.2186
26.6979 | -106.225
26.7016 | -106.2307
26.7053 | -106.2358
26.7091 | -106.2435
26.7128 | -106.2513
26.7166 | -106.2582
26.7203 | -106.2634
26.7241 | -106.2685
26.7278 | -106.2753
26.7315 | -106.2826
26.7353 | -106.2902
26.739 | -106.2979
26.7428 | -106.3063
26.7465 | -106.312
26.7503 | -106.3173
26.754 | -106.3245
26.7577 | -106.3332
26.7615 | -106.341
26.7652 | -106.3476
26.769 | -106.3563
26.7727 | -106.3659
26.7765 | -106.3747
26.7802 | -106.3841
26.7839 | -106.3917
26.7877 | -106.3995
26.7914 | -106.4072
26.7952 | -106.4141
26.7989 | -106.4215
26.8027 | -106.4273
26.8064 | -106.435
26.8101 | -106.4409
26.8139 | -106.4476
26.8176 | -106.4542
26.8214 | -106.4599
26.8251 | -106.4661
26.8289 | -106.4723
26.8326 | -106.4767
26.8363 | -106.4808
26.8401 | -106.4835
26.8438 | -106.4871
26.8476 | -106.4907
26.8513 | -106.4949
26.8551 | -106.4982
26.8588 | -106.501
26.8625 | -106.504
26.8663 | -106.5075
26.87 | -106.5096
26.8738 | -106.512
26.8775 | -106.5153
26.8813 | -106.517
26.885 | -106.5201
26.8887 | -106.5242
26.8925 | -106.5258
26.8962 | -106.5262
26.9 | -106.5263
26.9037 | -106.5279
26.9075 | -106.5296
26.9112 | -106.5297
26.9149 | -106.5277
26.9187 | -106.5253
26.9224 | -106.523
26.9262 | -106.5207
26.9299 | -106.5195
26.9337 | -106.5185
26.9374 | -106.5162
26.9411 | -106.5135
26.9449 | -106.5109
26.9486 | -106.5094
26.9524 | -106.5095
26.9561 | -106.5099
26.9599 | -106.508
26.9636 | -106.5064
26.9673 | -106.505
26.9711 | -106.502
26.9748 | -106.4982
26.9786 | -106.4943
26.9823 | -106.4897
26.9861 | -106.486
26.9898 | -106.4849
26.9935 | -106.4831
26.9973 | -106.481
27.001 | -106.4809
27.0048 | -106.4795
27.0085 | -106.4777
27.0123 | -106.4758
27.016 | -106.4738
27.0197 | -106.4723
27.0235 | -106.4704
27.0272 | -106.4707
27.031 | -106.472
27.0347 | -106.4714
27.0385 | -106.4704
27.0422 | -106.4707
27.0459 | -106.471
27.0497 | -106.4703
27.0534 | -106.4703
27.0572 | -106.4699
27.0609 | -106.4691
27.0647 | -106.4692
27.0684 | -106.4703
27.0721 | -106.4724
27.0759 | -106.4737
27.0796 | -106.4741
27.0834 | -106.4744
27.0871 | -106.4743
27.0909 | -106.4737
27.0946 | -106.4739
27.0983 | -106.4736
27.1021 | -106.4741
27.1058 | -106.4752
27.1096 | -106.4765
27.1133 | -106.4773
27.1171 | -106.4767
27.1208 | -106.4758
27.1245 | -106.4761
27.1283 | -106.4787
27.132 | -106.4804
27.1358 | -106.4796
27.1395 | -106.4775
27.1433 | -106.478
27.147 | -106.4793
27.1507 | -106.479
27.1545 | -106.4781
27.1582 | -106.4781
27.162 | -106.4779
27.1657 | -106.4783
27.1695 | -106.478
27.1732 | -106.4766
27.1769 | -106.4774
27.1807 | -106.4773
27.1844 | -106.4777
27.1882 | -106.4777
27.1919 | -106.4781
27.1957 | -106.4766
27.1994 | -106.4754
27.2031 | -106.4753
27.2069 | -106.4751
27.2106 | -106.4755
27.2144 | -106.4764
27.2181 | -106.4774
27.2219 | -106.4791
27.2256 | -106.4791
27.2293 | -106.4779
27.2331 | -106.4767
27.2368 | -106.4755
27.2406 | -106.474
27.2443 | -106.4707
27.2481 | -106.4662
27.2518 | -106.4633
27.2555 | -106.4604
27.2593 | -106.4577
27.263 | -106.4557
27.2668 | -106.4532
27.2705 | -106.4494
27.2743 | -106.4457
27.278 | -106.4432
27.2817 | -106.4393
27.2855 | -106.4371
27.2892 | -106.4344
27.293 | -106.4311
27.2967 | -106.4288
27.3005 | -106.4276
27.3042 | -106.4262
27.3079 | -106.4232
27.3117 | -106.4179
27.3154 | -106.4147
27.3192 | -106.412
27.3229 | -106.4074
27.3267 | -106.4032
27.3304 | -106.3996
27.3341 | -106.3961
27.3379 | -106.3915
27.3416 | -106.3884
27.3454 | -106.3858
27.3491 | -106.3815
27.3529 | -106.3794
27.3566 | -106.3761
27.3603 | -106.3736
27.3641 | -106.3724
27.3678 | -106.3715
27.3716 | -106.3672
27.3753 | -106.3645
27.3791 | -106.3631
27.3828 | -106.3602
27.3865 | -106.3557
27.3903 | -106.3522
27.394 | -106.3494
27.3978 | -106.3476
27.4015 | -106.3468
27.4053 | -106.3448
27.409 | -106.3413
27.4127 | -106.338
27.4165 | -106.3355
27.4202 | -106.3337
27.424 | -106.332
27.4277 | -106.3303
27.4315 | -106.3281
27.4352 | -106.3261
27.4389 | -106.324
27.4427 | -106.3222
27.4464 | -106.3204
27.4502 | -106.3182
27.4539 | -106.3151
27.4577 | -106.3108
27.4614 | -106.306
27.4651 | -106.302
27.4689 | -106.298
27.4726 | -106.2928
27.4764 | -106.2898
27.4801 | -106.2853
27.4839 | -106.2812
27.4876 | -106.2764
27.4913 | -106.2701
27.4951 | -106.2655
27.4988 | -106.2619
27.5026 | -106.2578
27.5063 | -106.2533
27.5101 | -106.2461
27.5138 | -106.2394
27.5175 | -106.2341
27.5213 | -106.2294
27.525 | -106.2259
27.5288 | -106.22
27.5325 | -106.2131
27.5363 | -106.2052
27.54 | -106.1983
27.5437 | -106.1918
27.5475 | -106.1853
27.5512 | -106.1783
27.555 | -106.1723
27.5587 | -106.1677
27.5625 | -106.1631
27.5662 | -106.1584
27.5699 | -106.1517
27.5737 | -106.147
27.5774 | -106.1419
27.5812 | -106.1367
27.5849 | -106.1325
27.5887 | -106.1295
27.5924 | -106.126
27.5961 | -106.1221
27.5999 | -106.1163
27.6036 | -106.1119
27.6074 | -106.1095
27.6111 | -106.107
27.6149 | -106.1048
27.6186 | -106.1018
27.6223 | -106.1
27.6261 | -106.1
27.6298 | -106.1003
27.6336 | -106.1011
27.6373 | -106.1014
27.6411 | -106.1013
27.6448 | -106.1006
27.6485 | -106.1001
27.6523 | -106.099
27.656 | -106.0995
27.6598 | -106.0998
27.6635 | -106.1005
27.6673 | -106.1007
27.671 | -106.1016
27.6747 | -106.1014
27.6785 | -106.1012
27.6822 | -106.1014
27.686 | -106.102
27.6897 | -106.1015
27.6935 | -106.0997
27.6972 | -106.0995
27.7009 | -106.101
27.7047 | -106.101
27.7084 | -106.1008
27.7122 | -106.1009
27.7159 | -106.1006
27.7197 | -106.0985
27.7234 | -106.0964
27.7271 | -106.0944
27.7309 | -106.0915
27.7346 | -106.0885
27.7384 | -106.0869
27.7421 | -106.0829
27.7459 | -106.0799
27.7496 | -106.0772
27.7533 | -106.0725
27.7571 | -106.0661
27.7608 | -106.0622
27.7646 | -106.058
27.7683 | -106.053
27.7721 | -106.0474
27.7758 | -106.043
27.7795 | -106.0349
27.7833 | -106.0266
27.787 | -106.0187
27.7908 | -106.0102
27.7945 | -106.002
27.7983 | -105.9954
27.802 | -105.9879
27.8057 | -105.9795
27.8095 | -105.9702
27.8132 | -105.9607
27.817 | -105.9503
27.8207 | -105.9396
27.8245 | -105.9297
27.8282 | -105.9182
27.8319 | -105.9075
27.8357 | -105.8973
27.8394 | -105.8886
27.8432 | -105.8785
27.8469 | -105.8685
27.8507 | -105.8571
27.8544 | -105.8454
27.8581 | -105.8365
27.8619 | -105.8296
27.8656 | -105.8215
27.8694 | -105.8116
27.8731 | -105.8023
27.8769 | -105.7964
27.8806 | -105.7903
27.8843 | -105.7824
27.8881 | -105.7745
27.8918 | -105.7672
27.8956 | -105.7597
27.8993 | -105.7537
27.9031 | -105.7479
27.9068 | -105.7417
27.9105 | -105.7366
27.9143 | -105.7315
27.918 | -105.7275
27.9218 | -105.724
27.9255 | -105.7214
27.9293 | -105.7184
27.933 | -105.7154
27.9367 | -105.7141
27.9405 | -105.7122
27.9442 | -105.7112
27.948 | -105.7099
27.9517 | -105.709
27.9555 | -105.7081
27.9592 | -105.706
27.9629 | -105.7054
27.9667 | -105.7061
27.9704 | -105.7057
27.9742 | -105.7056
27.9779 | -105.7056
27.9817 | -105.7069
27.9854 | -105.7072
27.9891 | -105.7084
27.9929 | -105.7123
27.9966 | -105.7152
28.0004 | -105.7171