## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 214 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 212
Hood | 02-92B/P | 151
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 999.5 | hPa |  | PRC PRO 01
Humidity | 43.5 | %/Hr
Internal temperature of calibration bench | 24.05 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1006 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.5 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.98 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3275 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 13/04/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 19/04/2018
Engineer | Anthony HUE | Approval |  | 23/04/2018
CEO | Robert GLEMOT | Approval |  | 23/04/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5068 | 57.4895
0.012209 | -36.0903 | 38.33
0.018313 | -35.0401 | 28.1811
0.024417 | -34.5752 | 21.9371
0.030521 | -34.4605 | 18.1534
0.036626 | -34.3155 | 15.1066
0.04273 | -34.2199 | 12.7559
0.048834 | -34.1737 | 11.2342
0.054938 | -34.1109 | 10.0968
0.061043 | -34.0525 | 9.0198
0.067147 | -34.042 | 8.205
0.073251 | -34.0423 | 7.4528
0.079355 | -34.0591 | 6.7968
0.08546 | -34.0719 | 6.2058
0.091564 | -34.0667 | 5.764
0.097668 | -34.0488 | 5.5085
0.10377 | -34.0274 | 5.2592
0.10988 | -34.0108 | 4.9819
0.11598 | -34.002 | 4.6911
0.12209 | -34.0037 | 4.4092
0.12819 | -34.013 | 4.1278
0.13429 | -34.0232 | 3.8426
0.1404 | -34.0262 | 3.5688
0.1465 | -34.0281 | 3.3147
0.15261 | -34.0257 | 3.097
0.15871 | -34.0198 | 2.9347
0.16482 | -34.0123 | 2.8292
0.17092 | -34.0099 | 2.7848
0.17702 | -34.0076 | 2.7094
0.18313 | -34.0065 | 2.6288
0.18923 | -34.0059 | 2.5196
0.19534 | -34.0045 | 2.3941
0.20144 | -34.0031 | 2.2545
0.20754 | -34.0004 | 2.136
0.21365 | -33.9994 | 2.0233
0.21975 | -33.999 | 1.9265
0.22586 | -33.9999 | 1.8398
0.23196 | -34.0022 | 1.7673
0.23807 | -34.006 | 1.7041
0.24417 | -34.0098 | 1.6453
0.25027 | -34.0135 | 1.583
0.25638 | -34.0157 | 1.5214
0.26248 | -34.0173 | 1.4555
0.26859 | -34.0181 | 1.3911
0.27469 | -34.0184 | 1.3211
0.2808 | -34.0193 | 1.2483
0.2869 | -34.02 | 1.1781
0.293 | -34.0208 | 1.1115
0.29911 | -34.0216 | 1.0512
0.30521 | -34.0224 | 0.99203
0.31132 | -34.0231 | 0.95225
0.31742 | -34.0242 | 0.90853
0.32353 | -34.0245 | 0.87252
0.32963 | -34.0253 | 0.84227
0.33573 | -34.0263 | 0.82058
0.34184 | -34.0276 | 0.79804
0.34794 | -34.0296 | 0.77388
0.35405 | -34.0316 | 0.74916
0.36015 | -34.034 | 0.72339
0.36626 | -34.0359 | 0.69571
0.37236 | -34.0379 | 0.65939
0.37846 | -34.04 | 0.61589
0.38457 | -34.0409 | 0.55997
0.39067 | -34.0419 | 0.51396
0.39678 | -34.0425 | 0.46297
0.40288 | -34.0429 | 0.4099
0.40899 | -34.0428 | 0.35876
0.41509 | -34.0422 | 0.30855
0.42119 | -34.0421 | 0.26558
0.4273 | -34.042 | 0.21982
0.4334 | -34.0419 | 0.17419
0.43951 | -34.0421 | 0.13129
0.44561 | -34.0429 | 0.091813
0.45172 | -34.0444 | 0.060095
0.45782 | -34.0451 | 0.023269
0.46392 | -34.0462 | -0.012724
0.47003 | -34.0472 | -0.045045
0.47613 | -34.0479 | -0.079029
0.48224 | -34.0484 | -0.11359
0.48834 | -34.0475 | -0.14802
0.49445 | -34.0473 | -0.18299
0.50055 | -34.0467 | -0.21488
0.50665 | -34.046 | -0.24879
0.51276 | -34.0451 | -0.28253
0.51886 | -34.0441 | -0.31617
0.52497 | -34.0429 | -0.34053
0.53107 | -34.0421 | -0.36992
0.53717 | -34.0415 | -0.39813
0.54328 | -34.0407 | -0.42124
0.54938 | -34.0405 | -0.44403
0.55549 | -34.0404 | -0.46533
0.56159 | -34.0415 | -0.47784
0.5677 | -34.0417 | -0.49318
0.5738 | -34.0427 | -0.51011
0.5799 | -34.0435 | -0.52579
0.58601 | -34.0445 | -0.53694
0.59211 | -34.0461 | -0.54443
0.59822 | -34.0474 | -0.55495
0.60432 | -34.0488 | -0.56788
0.61043 | -34.05 | -0.57696
0.61653 | -34.0512 | -0.5882
0.62263 | -34.0527 | -0.59856
0.62874 | -34.0534 | -0.61014
0.63484 | -34.0542 | -0.6204
0.64095 | -34.0547 | -0.63154
0.64705 | -34.0551 | -0.64227
0.65316 | -34.0553 | -0.65548
0.65926 | -34.0554 | -0.67179
0.66536 | -34.0543 | -0.68954
0.67147 | -34.054 | -0.70342
0.67757 | -34.0537 | -0.71828
0.68368 | -34.053 | -0.73747
0.68978 | -34.0522 | -0.75503
0.69589 | -34.0513 | -0.77429
0.70199 | -34.0504 | -0.79704
0.70809 | -34.0496 | -0.81647
0.7142 | -34.0489 | -0.83709
0.7203 | -34.0479 | -0.85652
0.72641 | -34.0472 | -0.87904
0.73251 | -34.0466 | -0.89919
0.73862 | -34.0468 | -0.92106
0.74472 | -34.0465 | -0.94124
0.75082 | -34.0462 | -0.9634
0.75693 | -34.0465 | -0.98532
0.76303 | -34.047 | -1.0074
0.76914 | -34.0488 | -1.028
0.77524 | -34.0497 | -1.0478
0.78135 | -34.0509 | -1.068
0.78745 | -34.0522 | -1.0882
0.79355 | -34.0533 | -1.1081
0.79966 | -34.055 | -1.1278
0.80576 | -34.0572 | -1.1476
0.81187 | -34.0587 | -1.1692
0.81797 | -34.0605 | -1.1905
0.82408 | -34.0621 | -1.2111
0.83018 | -34.0638 | -1.2317
0.83628 | -34.0656 | -1.2503
0.84239 | -34.0667 | -1.2709
0.84849 | -34.0681 | -1.2926
0.8546 | -34.0695 | -1.3125
0.8607 | -34.0707 | -1.3333
0.86681 | -34.0717 | -1.3523
0.87291 | -34.0727 | -1.3713
0.87901 | -34.0731 | -1.3895
0.88512 | -34.0738 | -1.4078
0.89122 | -34.0745 | -1.4262
0.89733 | -34.0749 | -1.4422
0.90343 | -34.0753 | -1.46
0.90953 | -34.0754 | -1.4748
0.91564 | -34.0755 | -1.491
0.92174 | -34.0757 | -1.5064
0.92785 | -34.0758 | -1.5215
0.93395 | -34.0758 | -1.5358
0.94006 | -34.076 | -1.5492
0.94616 | -34.0757 | -1.5617
0.95226 | -34.0758 | -1.5749
0.95837 | -34.0759 | -1.5892
0.96447 | -34.0761 | -1.6025
0.97058 | -34.0759 | -1.6164
0.97668 | -34.0757 | -1.6303
0.98279 | -34.0755 | -1.6435
0.98889 | -34.0753 | -1.6548
0.99499 | -34.0753 | -1.6687
1.0011 | -34.0751 | -1.6799
1.0072 | -34.0748 | -1.6921
1.0133 | -34.0745 | -1.7034
1.0194 | -34.0741 | -1.7167
1.0255 | -34.0738 | -1.7308
1.0316 | -34.0733 | -1.7432
1.0377 | -34.0729 | -1.7572
1.0438 | -34.0724 | -1.7703
1.0499 | -34.0717 | -1.7857
1.056 | -34.071 | -1.8004
1.0621 | -34.0704 | -1.8149
1.0682 | -34.07 | -1.8277
1.0743 | -34.0694 | -1.8433
1.0805 | -34.0689 | -1.8579
1.0866 | -34.0681 | -1.8731
1.0927 | -34.0673 | -1.8877
1.0988 | -34.0667 | -1.9028
1.1049 | -34.0659 | -1.9172
1.111 | -34.0652 | -1.9327
1.1171 | -34.0645 | -1.9474
1.1232 | -34.0641 | -1.9626
1.1293 | -34.0637 | -1.9768
1.1354 | -34.0634 | -1.9919
1.1415 | -34.0631 | -2.0075
1.1476 | -34.0625 | -2.023
1.1537 | -34.062 | -2.0378
1.1598 | -34.0621 | -2.0527
1.1659 | -34.0619 | -2.0661
1.172 | -34.0616 | -2.0802
1.1781 | -34.0614 | -2.0943
1.1842 | -34.0613 | -2.1087
1.1903 | -34.0616 | -2.12
1.1964 | -34.0615 | -2.1336
1.2025 | -34.0616 | -2.1477
1.2086 | -34.062 | -2.1603
1.2147 | -34.062 | -2.173
1.2209 | -34.0623 | -2.1845
1.227 | -34.063 | -2.1956
1.2331 | -34.0634 | -2.2079
1.2392 | -34.0637 | -2.2201
1.2453 | -34.0641 | -2.2315
1.2514 | -34.0645 | -2.2418
1.2575 | -34.0651 | -2.2532
1.2636 | -34.0659 | -2.2627
1.2697 | -34.0665 | -2.2724
1.2758 | -34.0671 | -2.2839
1.2819 | -34.0678 | -2.2943
1.288 | -34.0684 | -2.3055
1.2941 | -34.0692 | -2.3145
1.3002 | -34.0698 | -2.3231
1.3063 | -34.0703 | -2.3327
1.3124 | -34.071 | -2.3422
1.3185 | -34.0716 | -2.3518
1.3246 | -34.0723 | -2.3615
1.3307 | -34.0731 | -2.3706
1.3368 | -34.0738 | -2.3798
1.3429 | -34.0744 | -2.3893
1.349 | -34.0749 | -2.3991
1.3551 | -34.0756 | -2.409
1.3613 | -34.0761 | -2.4178
1.3674 | -34.0767 | -2.4283
1.3735 | -34.0771 | -2.4385
1.3796 | -34.0775 | -2.448
1.3857 | -34.0779 | -2.4578
1.3918 | -34.0784 | -2.4683
1.3979 | -34.0789 | -2.4779
1.404 | -34.0794 | -2.4896
1.4101 | -34.0798 | -2.5003
1.4162 | -34.0801 | -2.5112
1.4223 | -34.0806 | -2.5229
1.4284 | -34.081 | -2.5333
1.4345 | -34.0813 | -2.5451
1.4406 | -34.0816 | -2.5587
1.4467 | -34.082 | -2.5706
1.4528 | -34.0824 | -2.5822
1.4589 | -34.0829 | -2.5943
1.465 | -34.0831 | -2.6055
1.4711 | -34.0836 | -2.6171
1.4772 | -34.0841 | -2.6305
1.4833 | -34.0845 | -2.6437
1.4894 | -34.085 | -2.6561
1.4955 | -34.0855 | -2.6683
1.5016 | -34.086 | -2.681
1.5078 | -34.0864 | -2.695
1.5139 | -34.0868 | -2.7082
1.52 | -34.0872 | -2.721
1.5261 | -34.0875 | -2.7338
1.5322 | -34.0878 | -2.7479
1.5383 | -34.0882 | -2.7606
1.5444 | -34.0887 | -2.7734
1.5505 | -34.0891 | -2.787
1.5566 | -34.0896 | -2.8001
1.5627 | -34.0901 | -2.814
1.5688 | -34.0907 | -2.8278
1.5749 | -34.091 | -2.8405
1.581 | -34.0913 | -2.854
1.5871 | -34.0916 | -2.8678
1.5932 | -34.092 | -2.8809
1.5993 | -34.0924 | -2.8952
1.6054 | -34.0927 | -2.9086
1.6115 | -34.0931 | -2.9228
1.6176 | -34.0934 | -2.9356
1.6237 | -34.0936 | -2.9493
1.6298 | -34.0939 | -2.962
1.6359 | -34.0942 | -2.9749
1.642 | -34.0946 | -2.9871
1.6482 | -34.0948 | -3.0002
1.6543 | -34.0951 | -3.014
1.6604 | -34.0954 | -3.0263
1.6665 | -34.0955 | -3.0389
1.6726 | -34.0957 | -3.0516
1.6787 | -34.096 | -3.0643
1.6848 | -34.0959 | -3.0771
1.6909 | -34.0962 | -3.09
1.697 | -34.0964 | -3.1024
1.7031 | -34.0964 | -3.1152
1.7092 | -34.0964 | -3.1275
1.7153 | -34.0966 | -3.1397
1.7214 | -34.0964 | -3.1522
1.7275 | -34.0963 | -3.165
1.7336 | -34.0962 | -3.1781
1.7397 | -34.0964 | -3.1907
1.7458 | -34.0964 | -3.2031
1.7519 | -34.0964 | -3.2155
1.758 | -34.0964 | -3.2277
1.7641 | -34.0965 | -3.2396
1.7702 | -34.0965 | -3.2527
1.7763 | -34.0964 | -3.2644
1.7824 | -34.0964 | -3.2771
1.7885 | -34.0964 | -3.2898
1.7947 | -34.0965 | -3.3021
1.8008 | -34.0965 | -3.3136
1.8069 | -34.0965 | -3.3254
1.813 | -34.0966 | -3.3379
1.8191 | -34.0965 | -3.3508
1.8252 | -34.0966 | -3.3638
1.8313 | -34.0967 | -3.3763
1.8374 | -34.0968 | -3.3885
1.8435 | -34.0969 | -3.4007
1.8496 | -34.0969 | -3.414
1.8557 | -34.0971 | -3.427
1.8618 | -34.0973 | -3.4395
1.8679 | -34.0974 | -3.4518
1.874 | -34.0975 | -3.4641
1.8801 | -34.0977 | -3.4768
1.8862 | -34.0979 | -3.489
1.8923 | -34.098 | -3.5015
1.8984 | -34.0981 | -3.5139
1.9045 | -34.0982 | -3.5258
1.9106 | -34.0982 | -3.5384
1.9167 | -34.0983 | -3.5509
1.9228 | -34.0984 | -3.563
1.9289 | -34.0985 | -3.5753
1.9351 | -34.0986 | -3.5865
1.9412 | -34.0989 | -3.5981
1.9473 | -34.0991 | -3.6105
1.9534 | -34.0993 | -3.6223
1.9595 | -34.0995 | -3.6344
1.9656 | -34.0997 | -3.6471
1.9717 | -34.1 | -3.6592
1.9778 | -34.1002 | -3.6709
1.9839 | -34.1005 | -3.6828
1.99 | -34.1006 | -3.6943
1.9961 | -34.1007 | -3.7067
2.0022 | -34.101 | -3.7185
2.0083 | -34.101 | -3.7302
2.0144 | -34.1012 | -3.7419
2.0205 | -34.1015 | -3.754
2.0266 | -34.1019 | -3.7658
2.0327 | -34.1021 | -3.7778
2.0388 | -34.1025 | -3.7897
2.0449 | -34.1027 | -3.8015
2.051 | -34.1029 | -3.8132
2.0571 | -34.1032 | -3.8249
2.0632 | -34.1034 | -3.8368
2.0693 | -34.1036 | -3.8493
2.0754 | -34.1039 | -3.861
2.0816 | -34.1042 | -3.8734
2.0877 | -34.1044 | -3.8853
2.0938 | -34.1047 | -3.8977
2.0999 | -34.1049 | -3.9098
2.106 | -34.1051 | -3.9217
2.1121 | -34.1053 | -3.934
2.1182 | -34.1055 | -3.9452
2.1243 | -34.1058 | -3.9563
2.1304 | -34.106 | -3.9678
2.1365 | -34.1063 | -3.9799
2.1426 | -34.1065 | -3.9922
2.1487 | -34.1068 | -4.0037
2.1548 | -34.1072 | -4.0159
2.1609 | -34.1076 | -4.0276
2.167 | -34.108 | -4.0392
2.1731 | -34.1083 | -4.0515
2.1792 | -34.1085 | -4.0629
2.1853 | -34.1089 | -4.0752
2.1914 | -34.1092 | -4.0869
2.1975 | -34.1095 | -4.0987
2.2036 | -34.1097 | -4.1111
2.2097 | -34.11 | -4.123
2.2158 | -34.1104 | -4.134
2.222 | -34.1106 | -4.1448
2.2281 | -34.1109 | -4.1566
2.2342 | -34.1112 | -4.1683
2.2403 | -34.1114 | -4.1793
2.2464 | -34.1116 | -4.1903
2.2525 | -34.1119 | -4.202
2.2586 | -34.1122 | -4.213
2.2647 | -34.1125 | -4.2238
2.2708 | -34.1128 | -4.2354
2.2769 | -34.1131 | -4.2464
2.283 | -34.1133 | -4.2579
2.2891 | -34.1136 | -4.2692
2.2952 | -34.1137 | -4.2801
2.3013 | -34.114 | -4.2913
2.3074 | -34.1142 | -4.3022
2.3135 | -34.1144 | -4.3138
2.3196 | -34.1147 | -4.3244
2.3257 | -34.1149 | -4.3355
2.3318 | -34.115 | -4.347
2.3379 | -34.1152 | -4.3581
2.344 | -34.1153 | -4.3698
2.3501 | -34.1155 | -4.3809
2.3562 | -34.1157 | -4.3916
2.3623 | -34.1158 | -4.4024
2.3685 | -34.116 | -4.4133
2.3746 | -34.1161 | -4.4248
2.3807 | -34.1163 | -4.4357
2.3868 | -34.1165 | -4.4467
2.3929 | -34.1167 | -4.4578
2.399 | -34.1168 | -4.4686
2.4051 | -34.117 | -4.4795
2.4112 | -34.1172 | -4.4905
2.4173 | -34.1174 | -4.5015
2.4234 | -34.1175 | -4.5127
2.4295 | -34.1176 | -4.5238
2.4356 | -34.1178 | -4.535
2.4417 | -34.1179 | -4.5465
2.4478 | -34.118 | -4.5577
2.4539 | -34.1182 | -4.5683
2.46 | -34.1183 | -4.58
2.4661 | -34.1185 | -4.5914
2.4722 | -34.1186 | -4.602
2.4783 | -34.1187 | -4.6128
2.4844 | -34.1188 | -4.6242
2.4905 | -34.1189 | -4.6355
2.4966 | -34.119 | -4.6466
2.5027 | -34.1191 | -4.6574
2.5089 | -34.1193 | -4.6686
2.515 | -34.1193 | -4.6798
2.5211 | -34.1194 | -4.6909
2.5272 | -34.1195 | -4.7021
2.5333 | -34.1196 | -4.7134
2.5394 | -34.1197 | -4.7245
2.5455 | -34.1198 | -4.7362
2.5516 | -34.12 | -4.7469
2.5577 | -34.1202 | -4.7586
2.5638 | -34.1203 | -4.7706
2.5699 | -34.1204 | -4.7823
2.576 | -34.1204 | -4.7942
2.5821 | -34.1205 | -4.8057
2.5882 | -34.1206 | -4.8172
2.5943 | -34.1207 | -4.8283
2.6004 | -34.1209 | -4.8396
2.6065 | -34.1211 | -4.8504
2.6126 | -34.1212 | -4.8611
2.6187 | -34.1213 | -4.8719
2.6248 | -34.1214 | -4.883
2.6309 | -34.1216 | -4.8947
2.637 | -34.1218 | -4.9066
2.6431 | -34.122 | -4.9174
2.6492 | -34.1221 | -4.9291
2.6554 | -34.1223 | -4.9402
2.6615 | -34.1223 | -4.9515
2.6676 | -34.1226 | -4.9632
2.6737 | -34.1228 | -4.9747
2.6798 | -34.1231 | -4.9861
2.6859 | -34.1233 | -4.9976
2.692 | -34.1234 | -5.0088
2.6981 | -34.1236 | -5.0202
2.7042 | -34.1238 | -5.0317
2.7103 | -34.124 | -5.0427
2.7164 | -34.1243 | -5.0539
2.7225 | -34.1245 | -5.0656
2.7286 | -34.1247 | -5.0769
2.7347 | -34.1249 | -5.0877
2.7408 | -34.1252 | -5.0982
2.7469 | -34.1254 | -5.1098
2.753 | -34.1257 | -5.1207
2.7591 | -34.126 | -5.1323
2.7652 | -34.1262 | -5.144
2.7713 | -34.1263 | -5.1551
2.7774 | -34.1266 | -5.1663
2.7835 | -34.1268 | -5.1776
2.7896 | -34.127 | -5.1885
2.7958 | -34.1272 | -5.1995
2.8019 | -34.1275 | -5.2103
2.808 | -34.1278 | -5.2213
2.8141 | -34.1281 | -5.2328
2.8202 | -34.1284 | -5.244
2.8263 | -34.1287 | -5.2552
2.8324 | -34.129 | -5.2668
2.8385 | -34.1294 | -5.2779
2.8446 | -34.1297 | -5.289
2.8507 | -34.1301 | -5.3002
2.8568 | -34.1304 | -5.3109
2.8629 | -34.1307 | -5.3219
2.869 | -34.131 | -5.3333
2.8751 | -34.1313 | -5.3444
2.8812 | -34.1317 | -5.3559
2.8873 | -34.132 | -5.3671
2.8934 | -34.1323 | -5.3785
2.8995 | -34.1326 | -5.3899
2.9056 | -34.1329 | -5.4013
2.9117 | -34.1332 | -5.4127
2.9178 | -34.1336 | -5.4237
2.9239 | -34.134 | -5.4355
2.93 | -34.1344 | -5.4472
2.9361 | -34.1346 | -5.4584
2.9423 | -34.135 | -5.4694
2.9484 | -34.1354 | -5.4808
2.9545 | -34.1358 | -5.4921
2.9606 | -34.1362 | -5.5037
2.9667 | -34.1365 | -5.5152
2.9728 | -34.1368 | -5.5268
2.9789 | -34.1373 | -5.5382
2.985 | -34.1377 | -5.5498
2.9911 | -34.138 | -5.5614
2.9972 | -34.1385 | -5.5728
3.0033 | -34.1389 | -5.5846
3.0094 | -34.1393 | -5.596
3.0155 | -34.1397 | -5.6072
3.0216 | -34.14 | -5.6188
3.0277 | -34.1404 | -5.6308
3.0338 | -34.1408 | -5.6425
3.0399 | -34.1412 | -5.6541
3.046 | -34.1415 | -5.6659
3.0521 | -34.1418 | -5.6773
3.0582 | -34.1421 | -5.6897
3.0643 | -34.1424 | -5.7015
3.0704 | -34.1428 | -5.7136
3.0765 | -34.143 | -5.725
3.0827 | -34.1433 | -5.7365
3.0888 | -34.1437 | -5.7481
3.0949 | -34.1441 | -5.7598
3.101 | -34.1444 | -5.7716
3.1071 | -34.1447 | -5.7836
3.1132 | -34.1451 | -5.7957
3.1193 | -34.1455 | -5.8074
3.1254 | -34.1459 | -5.8191
3.1315 | -34.1462 | -5.8313
3.1376 | -34.1465 | -5.843
3.1437 | -34.1468 | -5.8549
3.1498 | -34.1472 | -5.8672
3.1559 | -34.1475 | -5.8789
3.162 | -34.1478 | -5.8913
3.1681 | -34.148 | -5.9036
3.1742 | -34.1484 | -5.9155
3.1803 | -34.1487 | -5.9274
3.1864 | -34.149 | -5.9391
3.1925 | -34.1493 | -5.9509
3.1986 | -34.1496 | -5.9632
3.2047 | -34.1499 | -5.9756
3.2108 | -34.1501 | -5.987
3.2169 | -34.1505 | -5.9991
3.223 | -34.1507 | -6.0111
3.2292 | -34.1509 | -6.0232
3.2353 | -34.1511 | -6.0344
3.2414 | -34.1514 | -6.0462
3.2475 | -34.1516 | -6.0579
3.2536 | -34.1519 | -6.0697
3.2597 | -34.1522 | -6.081
3.2658 | -34.1524 | -6.0923
3.2719 | -34.1526 | -6.1037
3.278 | -34.1529 | -6.1156
3.2841 | -34.1531 | -6.127
3.2902 | -34.1534 | -6.1388
3.2963 | -34.1536 | -6.1506
3.3024 | -34.1537 | -6.1621
3.3085 | -34.154 | -6.1739
3.3146 | -34.1542 | -6.1855
3.3207 | -34.1545 | -6.1973
3.3268 | -34.1547 | -6.2091
3.3329 | -34.1549 | -6.2202
3.339 | -34.155 | -6.2314
3.3451 | -34.1552 | -6.243
3.3512 | -34.1554 | -6.2541
3.3573 | -34.1556 | -6.2657
3.3634 | -34.1558 | -6.2772
3.3696 | -34.156 | -6.2889
3.3757 | -34.1562 | -6.3
3.3818 | -34.1564 | -6.3113
3.3879 | -34.1566 | -6.3227
3.394 | -34.1567 | -6.3344
3.4001 | -34.1569 | -6.3456
3.4062 | -34.157 | -6.3569
3.4123 | -34.1572 | -6.3679
3.4184 | -34.1574 | -6.3795
3.4245 | -34.1576 | -6.3907
3.4306 | -34.1577 | -6.4017
3.4367 | -34.1579 | -6.4127
3.4428 | -34.1581 | -6.4234
3.4489 | -34.1582 | -6.4347
3.455 | -34.1583 | -6.446
3.4611 | -34.1586 | -6.4575
3.4672 | -34.1587 | -6.4689
3.4733 | -34.1589 | -6.4804
3.4794 | -34.159 | -6.4915
3.4855 | -34.1592 | -6.5023
3.4916 | -34.1594 | -6.5131
3.4977 | -34.1595 | -6.5247
3.5038 | -34.1597 | -6.536
3.5099 | -34.1598 | -6.5467
3.5161 | -34.16 | -6.5574
3.5222 | -34.1601 | -6.5682
3.5283 | -34.1602 | -6.5793
3.5344 | -34.1603 | -6.5905
3.5405 | -34.1605 | -6.6014
3.5466 | -34.1607 | -6.6127
3.5527 | -34.1608 | -6.6238
3.5588 | -34.1609 | -6.6348
3.5649 | -34.161 | -6.6458
3.571 | -34.1611 | -6.6568
3.5771 | -34.1612 | -6.6683
3.5832 | -34.1614 | -6.6786
3.5893 | -34.1616 | -6.6896
3.5954 | -34.1618 | -6.7008
3.6015 | -34.1619 | -6.712
3.6076 | -34.1622 | -6.7229
3.6137 | -34.1623 | -6.7334
3.6198 | -34.1626 | -6.744
3.6259 | -34.1627 | -6.7553
3.632 | -34.1629 | -6.7663
3.6381 | -34.1631 | -6.7774
3.6442 | -34.1632 | -6.7883
3.6503 | -34.1634 | -6.7989
3.6565 | -34.1636 | -6.81
3.6626 | -34.1637 | -6.821
3.6687 | -34.164 | -6.8321
3.6748 | -34.1642 | -6.8429
3.6809 | -34.1645 | -6.8539
3.687 | -34.1646 | -6.8652
3.6931 | -34.1649 | -6.8758
3.6992 | -34.1651 | -6.8868
3.7053 | -34.1653 | -6.8976
3.7114 | -34.1655 | -6.9087
3.7175 | -34.1657 | -6.9197
3.7236 | -34.1659 | -6.9301
3.7297 | -34.1662 | -6.9413
3.7358 | -34.1664 | -6.9525
3.7419 | -34.1666 | -6.9637
3.748 | -34.1668 | -6.9747
3.7541 | -34.167 | -6.9856
3.7602 | -34.1673 | -6.9964
3.7663 | -34.1675 | -7.0073
3.7724 | -34.1677 | -7.0183
3.7785 | -34.168 | -7.029
3.7846 | -34.1682 | -7.0399
3.7907 | -34.1685 | -7.0511
3.7969 | -34.1688 | -7.0622
3.803 | -34.169 | -7.0728
3.8091 | -34.1693 | -7.0837
3.8152 | -34.1695 | -7.0945
3.8213 | -34.1698 | -7.1057
3.8274 | -34.17 | -7.1164
3.8335 | -34.1704 | -7.1275
3.8396 | -34.1707 | -7.1385
3.8457 | -34.171 | -7.1495
3.8518 | -34.1713 | -7.1604
3.8579 | -34.1716 | -7.1713
3.864 | -34.1719 | -7.1825
3.8701 | -34.1722 | -7.1938
3.8762 | -34.1725 | -7.2046
3.8823 | -34.1727 | -7.2157
3.8884 | -34.1731 | -7.2267
3.8945 | -34.1734 | -7.238
3.9006 | -34.1737 | -7.2493
3.9067 | -34.174 | -7.2607
3.9128 | -34.1743 | -7.2718
3.9189 | -34.1746 | -7.2826
3.925 | -34.1748 | -7.2937
3.9311 | -34.1751 | -7.3043
3.9372 | -34.1755 | -7.3155
3.9434 | -34.1757 | -7.3266
3.9495 | -34.1761 | -7.3377
3.9556 | -34.1764 | -7.3486
3.9617 | -34.1767 | -7.3593
3.9678 | -34.1771 | -7.37
3.9739 | -34.1773 | -7.381
3.98 | -34.1776 | -7.3921
3.9861 | -34.1779 | -7.403
3.9922 | -34.1782 | -7.4141
3.9983 | -34.1785 | -7.4247
4.0044 | -34.1788 | -7.4353
4.0105 | -34.1791 | -7.4463
4.0166 | -34.1794 | -7.4574
4.0227 | -34.1797 | -7.4683
4.0288 | -34.18 | -7.4798
4.0349 | -34.1802 | -7.4905
4.041 | -34.1805 | -7.5015
4.0471 | -34.1807 | -7.5126
4.0532 | -34.181 | -7.5232
4.0593 | -34.1814 | -7.5344
4.0654 | -34.1817 | -7.5451
4.0715 | -34.182 | -7.5566
4.0776 | -34.1823 | -7.5674
4.0838 | -34.1826 | -7.5788
4.0899 | -34.1829 | -7.5895
4.096 | -34.1832 | -7.6004
4.1021 | -34.1834 | -7.6112
4.1082 | -34.1837 | -7.6219
4.1143 | -34.184 | -7.6326
4.1204 | -34.1843 | -7.6433
4.1265 | -34.1846 | -7.6539
4.1326 | -34.1849 | -7.6645
4.1387 | -34.1852 | -7.6753
4.1448 | -34.1855 | -7.6863
4.1509 | -34.1857 | -7.6972
4.157 | -34.186 | -7.7085
4.1631 | -34.1863 | -7.7196
4.1692 | -34.1866 | -7.7303
4.1753 | -34.1869 | -7.7411
4.1814 | -34.1871 | -7.7517
4.1875 | -34.1874 | -7.7628
4.1936 | -34.1877 | -7.7738
4.1997 | -34.1879 | -7.7844
4.2058 | -34.1882 | -7.7955
4.2119 | -34.1884 | -7.8062
4.218 | -34.1887 | -7.8168
4.2241 | -34.1889 | -7.8275
4.2303 | -34.1892 | -7.8382
4.2364 | -34.1895 | -7.849
4.2425 | -34.1897 | -7.8598
4.2486 | -34.1899 | -7.8708
4.2547 | -34.1901 | -7.8814
4.2608 | -34.1903 | -7.8923
4.2669 | -34.1905 | -7.9028
4.273 | -34.1908 | -7.9138
4.2791 | -34.191 | -7.9246
4.2852 | -34.1913 | -7.9355
4.2913 | -34.1914 | -7.946
4.2974 | -34.1917 | -7.9571
4.3035 | -34.192 | -7.9681
4.3096 | -34.1922 | -7.9791
4.3157 | -34.1924 | -7.9897
4.3218 | -34.1926 | -8.0004
4.3279 | -34.1929 | -8.0112
4.334 | -34.1931 | -8.0219
4.3401 | -34.1933 | -8.0322
4.3462 | -34.1936 | -8.0431
4.3523 | -34.1939 | -8.0535
4.3584 | -34.194 | -8.0638
4.3645 | -34.1943 | -8.0745
4.3707 | -34.1945 | -8.0849
4.3768 | -34.1947 | -8.0956
4.3829 | -34.195 | -8.1066
4.389 | -34.1952 | -8.1173
4.3951 | -34.1954 | -8.1278
4.4012 | -34.1956 | -8.1388
4.4073 | -34.1958 | -8.1494
4.4134 | -34.196 | -8.16
4.4195 | -34.1963 | -8.1706
4.4256 | -34.1966 | -8.1811
4.4317 | -34.1968 | -8.1917
4.4378 | -34.197 | -8.202
4.4439 | -34.1972 | -8.2124
4.45 | -34.1974 | -8.223
4.4561 | -34.1975 | -8.2337
4.4622 | -34.1977 | -8.2444
4.4683 | -34.198 | -8.255
4.4744 | -34.1982 | -8.2659
4.4805 | -34.1984 | -8.2761
4.4866 | -34.1986 | -8.2867
4.4927 | -34.1989 | -8.2976
4.4988 | -34.1991 | -8.3082
4.5049 | -34.1993 | -8.3184
4.511 | -34.1996 | -8.3287
4.5172 | -34.1998 | -8.3393
4.5233 | -34.2 | -8.3503
4.5294 | -34.2002 | -8.3607
4.5355 | -34.2004 | -8.3713
4.5416 | -34.2006 | -8.3819
4.5477 | -34.2009 | -8.3925
4.5538 | -34.201 | -8.4026
4.5599 | -34.2013 | -8.4131
4.566 | -34.2015 | -8.4237
4.5721 | -34.2017 | -8.4343
4.5782 | -34.202 | -8.4451
4.5843 | -34.2022 | -8.4558
4.5904 | -34.2025 | -8.4664
4.5965 | -34.2027 | -8.4769
4.6026 | -34.2029 | -8.4879
4.6087 | -34.2031 | -8.4984
4.6148 | -34.2033 | -8.5093
4.6209 | -34.2036 | -8.5198
4.627 | -34.2039 | -8.5306
4.6331 | -34.204 | -8.5411
4.6392 | -34.2042 | -8.5518
4.6453 | -34.2045 | -8.5623
4.6514 | -34.2047 | -8.5729
4.6576 | -34.2049 | -8.5839
4.6637 | -34.2051 | -8.5941
4.6698 | -34.2053 | -8.6046
4.6759 | -34.2055 | -8.6155
4.682 | -34.2057 | -8.626
4.6881 | -34.2059 | -8.6366
4.6942 | -34.2062 | -8.6471
4.7003 | -34.2064 | -8.6578
4.7064 | -34.2066 | -8.6682
4.7125 | -34.2069 | -8.679
4.7186 | -34.207 | -8.6894
4.7247 | -34.2073 | -8.7
4.7308 | -34.2076 | -8.7106
4.7369 | -34.2078 | -8.7213
4.743 | -34.208 | -8.7322
4.7491 | -34.2082 | -8.7431
4.7552 | -34.2084 | -8.7538
4.7613 | -34.2086 | -8.7645
4.7674 | -34.2088 | -8.775
4.7735 | -34.2091 | -8.7861
4.7796 | -34.2093 | -8.7969
4.7857 | -34.2096 | -8.8077
4.7918 | -34.2098 | -8.8182
4.7979 | -34.2101 | -8.8287
4.8041 | -34.2103 | -8.8394
4.8102 | -34.2105 | -8.8497
4.8163 | -34.2108 | -8.8603
4.8224 | -34.211 | -8.871
4.8285 | -34.2113 | -8.8813
4.8346 | -34.2115 | -8.8917
4.8407 | -34.2117 | -8.9025
4.8468 | -34.212 | -8.9134
4.8529 | -34.2122 | -8.9243
4.859 | -34.2125 | -8.9352
4.8651 | -34.2127 | -8.9456
4.8712 | -34.213 | -8.9561
4.8773 | -34.2132 | -8.9668
4.8834 | -34.2134 | -8.9775
4.8895 | -34.2137 | -8.9879
4.8956 | -34.2139 | -8.9986
4.9017 | -34.2141 | -9.0091
4.9078 | -34.2143 | -9.02
4.9139 | -34.2145 | -9.0303
4.92 | -34.2148 | -9.0406
4.9261 | -34.2151 | -9.0509
4.9322 | -34.2154 | -9.0616
4.9383 | -34.2156 | -9.0723
4.9445 | -34.2159 | -9.0829
4.9506 | -34.2161 | -9.0934
4.9567 | -34.2164 | -9.1038
4.9628 | -34.2166 | -9.1143
4.9689 | -34.2168 | -9.1253
4.975 | -34.2171 | -9.1362
4.9811 | -34.2173 | -9.1466
4.9872 | -34.2176 | -9.1571
4.9933 | -34.2179 | -9.168
4.9994 | -34.2182 | -9.179
5.0055 | -34.2185 | -9.1897
5.0116 | -34.2187 | -9.2008
5.0177 | -34.2191 | -9.2117
5.0238 | -34.2193 | -9.2226
5.0299 | -34.2197 | -9.2336
5.036 | -34.2199 | -9.2443
5.0421 | -34.2202 | -9.2549
5.0482 | -34.2204 | -9.2656
5.0543 | -34.2207 | -9.276
5.0604 | -34.221 | -9.2867
5.0665 | -34.2213 | -9.2971
5.0726 | -34.2216 | -9.3081
5.0787 | -34.2219 | -9.319
5.0848 | -34.2222 | -9.3298
5.091 | -34.2224 | -9.3403
5.0971 | -34.2227 | -9.3506
5.1032 | -34.223 | -9.3617
5.1093 | -34.2233 | -9.3723
5.1154 | -34.2235 | -9.3828
5.1215 | -34.2238 | -9.3933
5.1276 | -34.224 | -9.4038
5.1337 | -34.2244 | -9.4142
5.1398 | -34.2246 | -9.4246
5.1459 | -34.2249 | -9.4355
5.152 | -34.2252 | -9.4464
5.1581 | -34.2255 | -9.4572
5.1642 | -34.2258 | -9.468
5.1703 | -34.2261 | -9.4787
5.1764 | -34.2264 | -9.4892
5.1825 | -34.2267 | -9.4997
5.1886 | -34.227 | -9.5101
5.1947 | -34.2272 | -9.5208
5.2008 | -34.2275 | -9.5316
5.2069 | -34.2278 | -9.5425
5.213 | -34.2281 | -9.553
5.2191 | -34.2284 | -9.5637
5.2252 | -34.2287 | -9.5745
5.2314 | -34.229 | -9.5853
5.2375 | -34.2293 | -9.5958
5.2436 | -34.2295 | -9.6066
5.2497 | -34.2298 | -9.617
5.2558 | -34.2302 | -9.6281
5.2619 | -34.2304 | -9.639
5.268 | -34.2308 | -9.6499
5.2741 | -34.2311 | -9.6606
5.2802 | -34.2314 | -9.6717
5.2863 | -34.2317 | -9.6824
5.2924 | -34.232 | -9.6933
5.2985 | -34.2323 | -9.7041
5.3046 | -34.2326 | -9.7148
5.3107 | -34.2329 | -9.7258
5.3168 | -34.2333 | -9.7366
5.3229 | -34.2335 | -9.7474
5.329 | -34.2338 | -9.7581
5.3351 | -34.2342 | -9.7686
5.3412 | -34.2344 | -9.7793
5.3473 | -34.2347 | -9.7901
5.3534 | -34.2351 | -9.801
5.3595 | -34.2354 | -9.8118
5.3656 | -34.2357 | -9.8225
5.3717 | -34.236 | -9.8336
5.3779 | -34.2363 | -9.8447
5.384 | -34.2366 | -9.8554
5.3901 | -34.2369 | -9.8665
5.3962 | -34.2372 | -9.8771
5.4023 | -34.2376 | -9.888
5.4084 | -34.2379 | -9.8988
5.4145 | -34.2382 | -9.9093
5.4206 | -34.2385 | -9.9201
5.4267 | -34.2388 | -9.9312
5.4328 | -34.2391 | -9.9421
5.4389 | -34.2395 | -9.9526
5.445 | -34.2397 | -9.9637
5.4511 | -34.2401 | -9.9743
5.4572 | -34.2404 | -9.9852
5.4633 | -34.2407 | -9.996
5.4694 | -34.241 | -10.0068
5.4755 | -34.2413 | -10.0174
5.4816 | -34.2416 | -10.0282
5.4877 | -34.242 | -10.0388
5.4938 | -34.2423 | -10.0494
5.4999 | -34.2426 | -10.0601
5.506 | -34.243 | -10.0708
5.5121 | -34.2432 | -10.0815
5.5183 | -34.2436 | -10.0922
5.5244 | -34.2439 | -10.1032
5.5305 | -34.2442 | -10.1139
5.5366 | -34.2445 | -10.1244
5.5427 | -34.2448 | -10.1354
5.5488 | -34.2451 | -10.1462
5.5549 | -34.2454 | -10.157
5.561 | -34.2457 | -10.1681
5.5671 | -34.2461 | -10.1787
5.5732 | -34.2464 | -10.1897
5.5793 | -34.2467 | -10.2008
5.5854 | -34.247 | -10.2116
5.5915 | -34.2473 | -10.2227
5.5976 | -34.2477 | -10.2335
5.6037 | -34.248 | -10.2444
5.6098 | -34.2483 | -10.2552
5.6159 | -34.2486 | -10.2658
5.622 | -34.2489 | -10.2767
5.6281 | -34.2492 | -10.2876
5.6342 | -34.2496 | -10.2987
5.6403 | -34.2499 | -10.3095
5.6464 | -34.2502 | -10.3205
5.6525 | -34.2505 | -10.3314
5.6586 | -34.2508 | -10.3419
5.6648 | -34.2512 | -10.3528
5.6709 | -34.2515 | -10.3638
5.677 | -34.2519 | -10.3746
5.6831 | -34.2522 | -10.3855
5.6892 | -34.2525 | -10.3965
5.6953 | -34.2528 | -10.4073
5.7014 | -34.2531 | -10.4183
5.7075 | -34.2534 | -10.4291
5.7136 | -34.2537 | -10.4398
5.7197 | -34.254 | -10.4507
5.7258 | -34.2543 | -10.4613
5.7319 | -34.2547 | -10.4722
5.738 | -34.255 | -10.483
5.7441 | -34.2554 | -10.4938
5.7502 | -34.2557 | -10.5043
5.7563 | -34.256 | -10.5151
5.7624 | -34.2564 | -10.5261
5.7685 | -34.2567 | -10.537
5.7746 | -34.257 | -10.5481
5.7807 | -34.2573 | -10.559
5.7868 | -34.2576 | -10.57
5.7929 | -34.2579 | -10.5807
5.799 | -34.2582 | -10.5914
5.8052 | -34.2585 | -10.6021
5.8113 | -34.2588 | -10.6128
5.8174 | -34.2591 | -10.6235
5.8235 | -34.2595 | -10.6343
5.8296 | -34.2598 | -10.6453
5.8357 | -34.2601 | -10.6561
5.8418 | -34.2604 | -10.667
5.8479 | -34.2607 | -10.6779
5.854 | -34.261 | -10.6889
5.8601 | -34.2613 | -10.6995
5.8662 | -34.2616 | -10.7105
5.8723 | -34.2619 | -10.7211
5.8784 | -34.2622 | -10.7319
5.8845 | -34.2625 | -10.7426
5.8906 | -34.2628 | -10.7534
5.8967 | -34.2632 | -10.7641
5.9028 | -34.2635 | -10.7747
5.9089 | -34.2639 | -10.7854
5.915 | -34.2642 | -10.796
5.9211 | -34.2645 | -10.8067
5.9272 | -34.2648 | -10.8173
5.9333 | -34.2651 | -10.8279
5.9394 | -34.2654 | -10.8387
5.9455 | -34.2658 | -10.8493
5.9517 | -34.2661 | -10.8599
5.9578 | -34.2665 | -10.8708
5.9639 | -34.2668 | -10.8812
5.97 | -34.2672 | -10.892
5.9761 | -34.2675 | -10.9027
5.9822 | -34.2678 | -10.9135
5.9883 | -34.2681 | -10.9242
5.9944 | -34.2685 | -10.9351
6.0005 | -34.2688 | -10.946
6.0066 | -34.2691 | -10.9565
6.0127 | -34.2695 | -10.967
6.0188 | -34.2698 | -10.978
6.0249 | -34.2701 | -10.9889
6.031 | -34.2704 | -10.9994
6.0371 | -34.2707 | -11.0102
6.0432 | -34.271 | -11.0208
6.0493 | -34.2714 | -11.0313
6.0554 | -34.2717 | -11.0419
6.0615 | -34.2721 | -11.0526
6.0676 | -34.2724 | -11.0635
6.0737 | -34.2728 | -11.0744
6.0798 | -34.273 | -11.085
6.0859 | -34.2734 | -11.0957
6.0921 | -34.2737 | -11.1062
6.0982 | -34.274 | -11.1169
6.1043 | -34.2744 | -11.1274
6.1104 | -34.2748 | -11.1383
6.1165 | -34.2751 | -11.1486
6.1226 | -34.2754 | -11.1592
6.1287 | -34.2758 | -11.1698
6.1348 | -34.2761 | -11.1805
6.1409 | -34.2764 | -11.1911
6.147 | -34.2767 | -11.2017
6.1531 | -34.2771 | -11.2123
6.1592 | -34.2774 | -11.2229
6.1653 | -34.2777 | -11.2334
6.1714 | -34.278 | -11.2441
6.1775 | -34.2784 | -11.2548
6.1836 | -34.2787 | -11.2655
6.1897 | -34.279 | -11.2761
6.1958 | -34.2793 | -11.287
6.2019 | -34.2797 | -11.2978
6.208 | -34.2801 | -11.3085
6.2141 | -34.2804 | -11.319
6.2202 | -34.2807 | -11.3298
6.2263 | -34.2811 | -11.3403
6.2325 | -34.2814 | -11.3506
6.2386 | -34.2817 | -11.361
6.2447 | -34.2821 | -11.3719
6.2508 | -34.2824 | -11.3825
6.2569 | -34.2828 | -11.393
6.263 | -34.2831 | -11.4036
6.2691 | -34.2834 | -11.4142
6.2752 | -34.2837 | -11.4249
6.2813 | -34.284 | -11.4354
6.2874 | -34.2843 | -11.4459
6.2935 | -34.2847 | -11.4565
6.2996 | -34.285 | -11.467
6.3057 | -34.2853 | -11.4777
6.3118 | -34.2856 | -11.4884
6.3179 | -34.2859 | -11.499
6.324 | -34.2863 | -11.5095
6.3301 | -34.2866 | -11.52
6.3362 | -34.287 | -11.5305
6.3423 | -34.2873 | -11.5413
6.3484 | -34.2876 | -11.5519
6.3545 | -34.2879 | -11.5622
6.3606 | -34.2883 | -11.5725
6.3667 | -34.2885 | -11.5829
6.3728 | -34.2889 | -11.5935
6.379 | -34.2892 | -11.6041
6.3851 | -34.2896 | -11.6144
6.3912 | -34.2899 | -11.6249
6.3973 | -34.2902 | -11.6355
6.4034 | -34.2905 | -11.6459
6.4095 | -34.2908 | -11.6564
6.4156 | -34.2911 | -11.667
6.4217 | -34.2915 | -11.6775
6.4278 | -34.2918 | -11.6881
6.4339 | -34.2921 | -11.6985
6.44 | -34.2925 | -11.709
6.4461 | -34.2928 | -11.7197
6.4522 | -34.2931 | -11.73
6.4583 | -34.2934 | -11.7407
6.4644 | -34.2937 | -11.7513
6.4705 | -34.2941 | -11.7615
6.4766 | -34.2944 | -11.772
6.4827 | -34.2947 | -11.7826
6.4888 | -34.295 | -11.7929
6.4949 | -34.2953 | -11.8032
6.501 | -34.2957 | -11.8137
6.5071 | -34.296 | -11.824
6.5132 | -34.2963 | -11.8345
6.5194 | -34.2966 | -11.8451
6.5255 | -34.297 | -11.8553
6.5316 | -34.2973 | -11.8659
6.5377 | -34.2976 | -11.8764
6.5438 | -34.2979 | -11.8867
6.5499 | -34.2982 | -11.897
6.556 | -34.2985 | -11.9075
6.5621 | -34.2989 | -11.9182
6.5682 | -34.2992 | -11.9287
6.5743 | -34.2996 | -11.939
6.5804 | -34.2999 | -11.9496
6.5865 | -34.3002 | -11.9601
6.5926 | -34.3005 | -11.9705
6.5987 | -34.3009 | -11.9806
6.6048 | -34.3012 | -11.9911
6.6109 | -34.3015 | -12.0015
6.617 | -34.3018 | -12.0121
6.6231 | -34.3021 | -12.0225
6.6292 | -34.3025 | -12.0329
6.6353 | -34.3028 | -12.0437
6.6414 | -34.3031 | -12.054
6.6475 | -34.3035 | -12.0644
6.6536 | -34.3038 | -12.0747
6.6597 | -34.3041 | -12.0849
6.6659 | -34.3044 | -12.0953
6.672 | -34.3048 | -12.1056
6.6781 | -34.3051 | -12.1161
6.6842 | -34.3054 | -12.1263
6.6903 | -34.3058 | -12.1367
6.6964 | -34.3061 | -12.1472
6.7025 | -34.3064 | -12.1574
6.7086 | -34.3067 | -12.1678
6.7147 | -34.3071 | -12.1783
6.7208 | -34.3074 | -12.1888
6.7269 | -34.3077 | -12.1991
6.733 | -34.308 | -12.2093
6.7391 | -34.3083 | -12.2197
6.7452 | -34.3087 | -12.2298
6.7513 | -34.309 | -12.2399
6.7574 | -34.3093 | -12.2504
6.7635 | -34.3097 | -12.2606
6.7696 | -34.31 | -12.2708
6.7757 | -34.3104 | -12.2812
6.7818 | -34.3107 | -12.2916
6.7879 | -34.3111 | -12.3018
6.794 | -34.3114 | -12.312
6.8001 | -34.3117 | -12.3222
6.8063 | -34.312 | -12.3326
6.8124 | -34.3123 | -12.343
6.8185 | -34.3126 | -12.3536
6.8246 | -34.3129 | -12.3637
6.8307 | -34.3133 | -12.3739
6.8368 | -34.3136 | -12.3845
6.8429 | -34.3139 | -12.3948
6.849 | -34.3143 | -12.4052
6.8551 | -34.3147 | -12.4154
6.8612 | -34.315 | -12.4257
6.8673 | -34.3153 | -12.436
6.8734 | -34.3156 | -12.4464
6.8795 | -34.3159 | -12.4568
6.8856 | -34.3163 | -12.4671
6.8917 | -34.3166 | -12.4773
6.8978 | -34.3169 | -12.4874
6.9039 | -34.3172 | -12.4974
6.91 | -34.3176 | -12.5078
6.9161 | -34.3179 | -12.5178
6.9222 | -34.3183 | -12.5281
6.9283 | -34.3186 | -12.5381
6.9344 | -34.319 | -12.5485
6.9405 | -34.3193 | -12.5587
6.9466 | -34.3197 | -12.5692
6.9528 | -34.32 | -12.5796
6.9589 | -34.3204 | -12.59
6.965 | -34.3207 | -12.6004
6.9711 | -34.3211 | -12.6107
6.9772 | -34.3215 | -12.621
6.9833 | -34.3218 | -12.6313
6.9894 | -34.3221 | -12.6418
6.9955 | -34.3225 | -12.6518
7.0016 | -34.3228 | -12.662
7.0077 | -34.3232 | -12.6723
7.0138 | -34.3235 | -12.6825
7.0199 | -34.3238 | -12.6926
7.026 | -34.3242 | -12.7028
7.0321 | -34.3245 | -12.7129
7.0382 | -34.3249 | -12.7228
7.0443 | -34.3253 | -12.7332
7.0504 | -34.3256 | -12.7436
7.0565 | -34.3259 | -12.7539
7.0626 | -34.3263 | -12.7639
7.0687 | -34.3266 | -12.7737
7.0748 | -34.327 | -12.7837
7.0809 | -34.3273 | -12.7939
7.087 | -34.3277 | -12.8041
7.0932 | -34.328 | -12.8144
7.0993 | -34.3283 | -12.8246
7.1054 | -34.3287 | -12.8346
7.1115 | -34.3291 | -12.8446
7.1176 | -34.3294 | -12.8548
7.1237 | -34.3297 | -12.8649
7.1298 | -34.3301 | -12.875
7.1359 | -34.3304 | -12.8852
7.142 | -34.3308 | -12.8952
7.1481 | -34.3311 | -12.9053
7.1542 | -34.3314 | -12.9155
7.1603 | -34.3318 | -12.9256
7.1664 | -34.3322 | -12.9358
7.1725 | -34.3325 | -12.946
7.1786 | -34.3329 | -12.9562
7.1847 | -34.3332 | -12.9664
7.1908 | -34.3336 | -12.9767
7.1969 | -34.3339 | -12.987
7.203 | -34.3343 | -12.9971
7.2091 | -34.3346 | -13.0071
7.2152 | -34.335 | -13.0173
7.2213 | -34.3353 | -13.0273
7.2274 | -34.3357 | -13.0375
7.2335 | -34.336 | -13.0476
7.2397 | -34.3364 | -13.0577
7.2458 | -34.3368 | -13.0676
7.2519 | -34.3371 | -13.0778
7.258 | -34.3375 | -13.0881
7.2641 | -34.3378 | -13.0982
7.2702 | -34.3382 | -13.1086
7.2763 | -34.3386 | -13.1187
7.2824 | -34.3389 | -13.129
7.2885 | -34.3393 | -13.1392
7.2946 | -34.3396 | -13.1494
7.3007 | -34.34 | -13.1596
7.3068 | -34.3403 | -13.17
7.3129 | -34.3407 | -13.1804
7.319 | -34.341 | -13.1905
7.3251 | -34.3414 | -13.2005
7.3312 | -34.3417 | -13.2107
7.3373 | -34.3421 | -13.2206
7.3434 | -34.3425 | -13.2307
7.3495 | -34.3429 | -13.2407
7.3556 | -34.3432 | -13.251
7.3617 | -34.3436 | -13.2613
7.3678 | -34.3439 | -13.2713
7.3739 | -34.3443 | -13.2817
7.3801 | -34.3446 | -13.2921
7.3862 | -34.345 | -13.3022
7.3923 | -34.3454 | -13.3125
7.3984 | -34.3457 | -13.3224
7.4045 | -34.3461 | -13.3325
7.4106 | -34.3465 | -13.3426
7.4167 | -34.3469 | -13.3529
7.4228 | -34.3473 | -13.3631
7.4289 | -34.3476 | -13.3733
7.435 | -34.348 | -13.3832
7.4411 | -34.3484 | -13.3935
7.4472 | -34.3487 | -13.4037
7.4533 | -34.3491 | -13.4137
7.4594 | -34.3495 | -13.424
7.4655 | -34.3498 | -13.434
7.4716 | -34.3502 | -13.4443
7.4777 | -34.3505 | -13.4545
7.4838 | -34.3509 | -13.4645
7.4899 | -34.3513 | -13.4746
7.496 | -34.3516 | -13.4846
7.5021 | -34.352 | -13.4947
7.5082 | -34.3524 | -13.5049
7.5143 | -34.3527 | -13.515
7.5204 | -34.3531 | -13.5253
7.5266 | -34.3535 | -13.5354
7.5327 | -34.3538 | -13.5454
7.5388 | -34.3542 | -13.5559
7.5449 | -34.3546 | -13.5662
7.551 | -34.355 | -13.5762
7.5571 | -34.3553 | -13.5865
7.5632 | -34.3557 | -13.5967
7.5693 | -34.3561 | -13.6068
7.5754 | -34.3565 | -13.6168
7.5815 | -34.3568 | -13.627
7.5876 | -34.3572 | -13.6374
7.5937 | -34.3576 | -13.6477
7.5998 | -34.358 | -13.6578
7.6059 | -34.3584 | -13.6681
7.612 | -34.3588 | -13.6783
7.6181 | -34.3591 | -13.6884
7.6242 | -34.3595 | -13.6988
7.6303 | -34.3599 | -13.7092
7.6364 | -34.3602 | -13.7194
7.6425 | -34.3606 | -13.7296
7.6486 | -34.361 | -13.7398
7.6547 | -34.3613 | -13.7499
7.6608 | -34.3617 | -13.7601
7.667 | -34.3621 | -13.7704
7.6731 | -34.3625 | -13.7806
7.6792 | -34.3629 | -13.791
7.6853 | -34.3632 | -13.8013
7.6914 | -34.3636 | -13.8116
7.6975 | -34.364 | -13.8219
7.7036 | -34.3644 | -13.8322
7.7097 | -34.3648 | -13.8425
7.7158 | -34.3652 | -13.8528
7.7219 | -34.3655 | -13.8629
7.728 | -34.3659 | -13.8732
7.7341 | -34.3663 | -13.8835
7.7402 | -34.3667 | -13.8938
7.7463 | -34.3671 | -13.9041
7.7524 | -34.3674 | -13.9145
7.7585 | -34.3678 | -13.9248
7.7646 | -34.3682 | -13.935
7.7707 | -34.3685 | -13.9455
7.7768 | -34.3689 | -13.9558
7.7829 | -34.3693 | -13.9663
7.789 | -34.3696 | -13.9768
7.7951 | -34.37 | -13.987
7.8012 | -34.3704 | -13.9973
7.8073 | -34.3708 | -14.0078
7.8135 | -34.3712 | -14.0181
7.8196 | -34.3716 | -14.0285
7.8257 | -34.372 | -14.039
7.8318 | -34.3723 | -14.0494
7.8379 | -34.3727 | -14.0595
7.844 | -34.3731 | -14.0698
7.8501 | -34.3735 | -14.0801
7.8562 | -34.3739 | -14.0906
7.8623 | -34.3743 | -14.1008
7.8684 | -34.3746 | -14.1112
7.8745 | -34.375 | -14.1215
7.8806 | -34.3754 | -14.132
7.8867 | -34.3758 | -14.1421
7.8928 | -34.3762 | -14.1522
7.8989 | -34.3766 | -14.1625
7.905 | -34.3769 | -14.1729
7.9111 | -34.3773 | -14.1832
7.9172 | -34.3777 | -14.1935
7.9233 | -34.3781 | -14.2039
7.9294 | -34.3785 | -14.2143
7.9355 | -34.3788 | -14.2244
7.9416 | -34.3792 | -14.2348
7.9477 | -34.3796 | -14.2449
7.9539 | -34.38 | -14.2552
7.96 | -34.3803 | -14.2656
7.9661 | -34.3807 | -14.276
7.9722 | -34.3812 | -14.2862
7.9783 | -34.3815 | -14.2967
7.9844 | -34.382 | -14.307
7.9905 | -34.3823 | -14.3175
7.9966 | -34.3827 | -14.3279
8.0027 | -34.3831 | -14.3382
8.0088 | -34.3834 | -14.3485
8.0149 | -34.3838 | -14.3589
8.021 | -34.3842 | -14.3693
8.0271 | -34.3846 | -14.3799
8.0332 | -34.385 | -14.3901
8.0393 | -34.3854 | -14.4006
8.0454 | -34.3858 | -14.4112
8.0515 | -34.3862 | -14.4215
8.0576 | -34.3866 | -14.4321
8.0637 | -34.3869 | -14.4423
8.0698 | -34.3873 | -14.4525
8.0759 | -34.3876 | -14.4628
8.082 | -34.388 | -14.473
8.0881 | -34.3884 | -14.4834
8.0942 | -34.3888 | -14.4938
8.1004 | -34.3893 | -14.5043
8.1065 | -34.3896 | -14.5145
8.1126 | -34.39 | -14.5248
8.1187 | -34.3904 | -14.5352
8.1248 | -34.3907 | -14.5454
8.1309 | -34.3911 | -14.5556
8.137 | -34.3915 | -14.5658
8.1431 | -34.3919 | -14.5763
8.1492 | -34.3923 | -14.5866
8.1553 | -34.3927 | -14.597
8.1614 | -34.3931 | -14.6075
8.1675 | -34.3934 | -14.6178
8.1736 | -34.3938 | -14.6283
8.1797 | -34.3942 | -14.6387
8.1858 | -34.3946 | -14.6492
8.1919 | -34.395 | -14.6595
8.198 | -34.3953 | -14.6696
8.2041 | -34.3957 | -14.68
8.2102 | -34.3961 | -14.6904
8.2163 | -34.3965 | -14.7005
8.2224 | -34.3969 | -14.7108
8.2285 | -34.3973 | -14.721
8.2346 | -34.3977 | -14.7312
8.2408 | -34.3981 | -14.7414
8.2469 | -34.3985 | -14.7518
8.253 | -34.3989 | -14.762
8.2591 | -34.3993 | -14.7724
8.2652 | -34.3997 | -14.7827
8.2713 | -34.4001 | -14.7931
8.2774 | -34.4005 | -14.8035
8.2835 | -34.4009 | -14.8141
8.2896 | -34.4012 | -14.8243
8.2957 | -34.4016 | -14.8347
8.3018 | -34.402 | -14.845
8.3079 | -34.4024 | -14.8551
8.314 | -34.4028 | -14.8655
8.3201 | -34.4033 | -14.8757
8.3262 | -34.4036 | -14.8861
8.3323 | -34.404 | -14.8963
8.3384 | -34.4044 | -14.9068
8.3445 | -34.4048 | -14.9171
8.3506 | -34.4052 | -14.9275
8.3567 | -34.4056 | -14.9379
8.3628 | -34.406 | -14.9482
8.3689 | -34.4064 | -14.9583
8.375 | -34.4068 | -14.9688
8.3812 | -34.4071 | -14.979
8.3873 | -34.4075 | -14.9893
8.3934 | -34.4079 | -14.9997
8.3995 | -34.4083 | -15.0099
8.4056 | -34.4087 | -15.0202
8.4117 | -34.4091 | -15.0306
8.4178 | -34.4095 | -15.0409
8.4239 | -34.4098 | -15.0511
8.43 | -34.4102 | -15.0614
8.4361 | -34.4106 | -15.0718
8.4422 | -34.411 | -15.082
8.4483 | -34.4114 | -15.0922
8.4544 | -34.4118 | -15.1025
8.4605 | -34.4122 | -15.1127
8.4666 | -34.4126 | -15.1232
8.4727 | -34.4129 | -15.1335
8.4788 | -34.4133 | -15.1438
8.4849 | -34.4137 | -15.154
8.491 | -34.4141 | -15.1643
8.4971 | -34.4145 | -15.1744
8.5032 | -34.4149 | -15.1845
8.5093 | -34.4153 | -15.1947
8.5154 | -34.4157 | -15.2051
8.5215 | -34.4161 | -15.2152
8.5277 | -34.4164 | -15.2255
8.5338 | -34.4168 | -15.2357
8.5399 | -34.4172 | -15.246
8.546 | -34.4176 | -15.256
8.5521 | -34.418 | -15.2663
8.5582 | -34.4184 | -15.2765
8.5643 | -34.4188 | -15.2865
8.5704 | -34.4192 | -15.2965
8.5765 | -34.4196 | -15.3069
8.5826 | -34.42 | -15.3169
8.5887 | -34.4204 | -15.327
8.5948 | -34.4208 | -15.3373
8.6009 | -34.4212 | -15.3474
8.607 | -34.4216 | -15.3577
8.6131 | -34.422 | -15.368
8.6192 | -34.4224 | -15.378
8.6253 | -34.4228 | -15.3882
8.6314 | -34.4232 | -15.3983
8.6375 | -34.4235 | -15.4086
8.6436 | -34.4239 | -15.4188
8.6497 | -34.4243 | -15.4288
8.6558 | -34.4247 | -15.439
8.6619 | -34.425 | -15.4492
8.6681 | -34.4254 | -15.4596
8.6742 | -34.4259 | -15.4697
8.6803 | -34.4262 | -15.4799
8.6864 | -34.4266 | -15.4898
8.6925 | -34.427 | -15.4998
8.6986 | -34.4274 | -15.51
8.7047 | -34.4278 | -15.5204
8.7108 | -34.4282 | -15.5305
8.7169 | -34.4286 | -15.5407
8.723 | -34.429 | -15.5507
8.7291 | -34.4294 | -15.5607
8.7352 | -34.4298 | -15.571
8.7413 | -34.4301 | -15.5812
8.7474 | -34.4305 | -15.5912
8.7535 | -34.4309 | -15.6014
8.7596 | -34.4314 | -15.6115
8.7657 | -34.4317 | -15.6218
8.7718 | -34.4321 | -15.632
8.7779 | -34.4325 | -15.6422
8.784 | -34.4329 | -15.6524
8.7901 | -34.4333 | -15.6625
8.7962 | -34.4337 | -15.6724
8.8023 | -34.4341 | -15.6825
8.8084 | -34.4345 | -15.6926
8.8146 | -34.4349 | -15.7027
8.8207 | -34.4353 | -15.7128
8.8268 | -34.4357 | -15.7225
8.8329 | -34.4361 | -15.7326
8.839 | -34.4365 | -15.7428
8.8451 | -34.4369 | -15.7528
8.8512 | -34.4373 | -15.7628
8.8573 | -34.4377 | -15.773
8.8634 | -34.4381 | -15.7828
8.8695 | -34.4385 | -15.7931
8.8756 | -34.4388 | -15.8032
8.8817 | -34.4392 | -15.8131
8.8878 | -34.4396 | -15.8234
8.8939 | -34.44 | -15.8335
8.9 | -34.4404 | -15.8436
8.9061 | -34.4408 | -15.8537
8.9122 | -34.4412 | -15.8637
8.9183 | -34.4416 | -15.8739
8.9244 | -34.442 | -15.8841
8.9305 | -34.4425 | -15.8939
8.9366 | -34.4429 | -15.904
8.9427 | -34.4433 | -15.9141
8.9488 | -34.4437 | -15.9241
8.955 | -34.4441 | -15.9341
8.9611 | -34.4445 | -15.9443
8.9672 | -34.4449 | -15.9543
8.9733 | -34.4453 | -15.9642
8.9794 | -34.4457 | -15.9741
8.9855 | -34.4461 | -15.9843
8.9916 | -34.4465 | -15.9945
8.9977 | -34.4469 | -16.0045
9.0038 | -34.4474 | -16.0144
9.0099 | -34.4478 | -16.0245
9.016 | -34.4482 | -16.0346
9.0221 | -34.4486 | -16.0445
9.0282 | -34.449 | -16.0544
9.0343 | -34.4494 | -16.0645
9.0404 | -34.4499 | -16.0745
9.0465 | -34.4503 | -16.0847
9.0526 | -34.4507 | -16.0946
9.0587 | -34.4511 | -16.1047
9.0648 | -34.4516 | -16.1146
9.0709 | -34.4519 | -16.1246
9.077 | -34.4524 | -16.1347
9.0831 | -34.4528 | -16.1449
9.0892 | -34.4532 | -16.155
9.0953 | -34.4536 | -16.165
9.1015 | -34.4541 | -16.175
9.1076 | -34.4545 | -16.1849
9.1137 | -34.4549 | -16.1949
9.1198 | -34.4553 | -16.2049
9.1259 | -34.4557 | -16.2149
9.132 | -34.4562 | -16.225
9.1381 | -34.4566 | -16.235
9.1442 | -34.457 | -16.2446
9.1503 | -34.4574 | -16.2547
9.1564 | -34.4579 | -16.2647
9.1625 | -34.4583 | -16.2747
9.1686 | -34.4587 | -16.2846
9.1747 | -34.4591 | -16.2946
9.1808 | -34.4596 | -16.3045
9.1869 | -34.46 | -16.3145
9.193 | -34.4604 | -16.3245
9.1991 | -34.4608 | -16.3346
9.2052 | -34.4613 | -16.3446
9.2113 | -34.4617 | -16.3545
9.2174 | -34.4622 | -16.3645
9.2235 | -34.4626 | -16.3744
9.2296 | -34.463 | -16.3844
9.2357 | -34.4634 | -16.3945
9.2419 | -34.4639 | -16.4044
9.248 | -34.4644 | -16.4143
9.2541 | -34.4648 | -16.4244
9.2602 | -34.4652 | -16.4345
9.2663 | -34.4657 | -16.4442
9.2724 | -34.4661 | -16.4541
9.2785 | -34.4665 | -16.464
9.2846 | -34.467 | -16.4738
9.2907 | -34.4674 | -16.4838
9.2968 | -34.4678 | -16.4938
9.3029 | -34.4682 | -16.504
9.309 | -34.4687 | -16.514
9.3151 | -34.4691 | -16.5238
9.3212 | -34.4695 | -16.5336
9.3273 | -34.47 | -16.5435
9.3334 | -34.4704 | -16.5533
9.3395 | -34.4708 | -16.5631
9.3456 | -34.4712 | -16.5729
9.3517 | -34.4716 | -16.5829
9.3578 | -34.4721 | -16.5928
9.3639 | -34.4725 | -16.6026
9.37 | -34.473 | -16.6126
9.3761 | -34.4734 | -16.6227
9.3822 | -34.4739 | -16.6327
9.3884 | -34.4743 | -16.6423
9.3945 | -34.4747 | -16.6523
9.4006 | -34.4752 | -16.6623
9.4067 | -34.4756 | -16.6723
9.4128 | -34.4761 | -16.6823
9.4189 | -34.4765 | -16.6924
9.425 | -34.477 | -16.7022
9.4311 | -34.4774 | -16.7121
9.4372 | -34.4779 | -16.722
9.4433 | -34.4783 | -16.7319
9.4494 | -34.4787 | -16.7416
9.4555 | -34.4792 | -16.7517
9.4616 | -34.4797 | -16.7615
9.4677 | -34.4801 | -16.7714
9.4738 | -34.4805 | -16.7813
9.4799 | -34.481 | -16.7912
9.486 | -34.4814 | -16.8012
9.4921 | -34.4818 | -16.8111
9.4982 | -34.4823 | -16.8208
9.5043 | -34.4828 | -16.8307
9.5104 | -34.4832 | -16.8404
9.5165 | -34.4836 | -16.8503
9.5226 | -34.4841 | -16.8603
9.5288 | -34.4845 | -16.87
9.5349 | -34.4849 | -16.8799
9.541 | -34.4854 | -16.8897
9.5471 | -34.4858 | -16.8995
9.5532 | -34.4863 | -16.9095
9.5593 | -34.4867 | -16.9194
9.5654 | -34.4872 | -16.9292
9.5715 | -34.4876 | -16.9391
9.5776 | -34.488 | -16.9491
9.5837 | -34.4885 | -16.959
9.5898 | -34.4889 | -16.969
9.5959 | -34.4893 | -16.9789
9.602 | -34.4898 | -16.9886
9.6081 | -34.4902 | -16.9985
9.6142 | -34.4907 | -17.0083
9.6203 | -34.4911 | -17.0182
9.6264 | -34.4915 | -17.028
9.6325 | -34.492 | -17.0379
9.6386 | -34.4924 | -17.0476
9.6447 | -34.4929 | -17.0575
9.6508 | -34.4933 | -17.0673
9.6569 | -34.4937 | -17.0772
9.663 | -34.4942 | -17.0871
9.6691 | -34.4947 | -17.097
9.6753 | -34.4951 | -17.1068
9.6814 | -34.4956 | -17.1166
9.6875 | -34.496 | -17.1262
9.6936 | -34.4964 | -17.136
9.6997 | -34.4969 | -17.146
9.7058 | -34.4974 | -17.1557
9.7119 | -34.4978 | -17.1657
9.718 | -34.4983 | -17.1757
9.7241 | -34.4987 | -17.1854
9.7302 | -34.4992 | -17.1954
9.7363 | -34.4996 | -17.2051
9.7424 | -34.5001 | -17.2149
9.7485 | -34.5005 | -17.2248
9.7546 | -34.501 | -17.2345
9.7607 | -34.5014 | -17.2446
9.7668 | -34.5019 | -17.2545
9.7729 | -34.5023 | -17.2643
9.779 | -34.5028 | -17.2739
9.7851 | -34.5033 | -17.2839
9.7912 | -34.5037 | -17.2937
9.7973 | -34.5042 | -17.3037
9.8034 | -34.5047 | -17.3133
9.8095 | -34.5051 | -17.3233
9.8157 | -34.5056 | -17.3332
9.8218 | -34.506 | -17.3432
9.8279 | -34.5065 | -17.3529
9.834 | -34.5069 | -17.3628
9.8401 | -34.5073 | -17.3726
9.8462 | -34.5078 | -17.3822
9.8523 | -34.5083 | -17.3921
9.8584 | -34.5087 | -17.4019
9.8645 | -34.5091 | -17.4118
9.8706 | -34.5096 | -17.4217
9.8767 | -34.51 | -17.4314
9.8828 | -34.5104 | -17.4412
9.8889 | -34.5109 | -17.4508
9.895 | -34.5113 | -17.4607
9.9011 | -34.5118 | -17.4705
9.9072 | -34.5122 | -17.4803
9.9133 | -34.5127 | -17.4901
9.9194 | -34.5131 | -17.4998
9.9255 | -34.5136 | -17.5096
9.9316 | -34.5141 | -17.5193
9.9377 | -34.5145 | -17.5291
9.9438 | -34.515 | -17.539
9.9499 | -34.5154 | -17.549
9.956 | -34.5159 | -17.5588
9.9622 | -34.5163 | -17.5687
9.9683 | -34.5168 | -17.5785
9.9744 | -34.5172 | -17.5884
9.9805 | -34.5177 | -17.5982
9.9866 | -34.5181 | -17.6079
9.9927 | -34.5185 | -17.6176
9.9988 | -34.519 | -17.6275
10.0049 | -34.5195 | -17.6373
10.011 | -34.5199 | -17.6472
10.0171 | -34.5204 | -17.6569
10.0232 | -34.5208 | -17.6666
10.0293 | -34.5213 | -17.6765
10.0354 | -34.5218 | -17.6862
10.0415 | -34.5222 | -17.6961
10.0476 | -34.5227 | -17.706
10.0537 | -34.5231 | -17.7159
10.0598 | -34.5236 | -17.7257
10.0659 | -34.524 | -17.7355
10.072 | -34.5245 | -17.7451
10.0781 | -34.5249 | -17.7549
10.0842 | -34.5254 | -17.7646
10.0903 | -34.5258 | -17.7744
10.0964 | -34.5263 | -17.7839
10.1026 | -34.5267 | -17.7938
10.1087 | -34.5272 | -17.8036
10.1148 | -34.5276 | -17.8134
10.1209 | -34.5281 | -17.8231
10.127 | -34.5285 | -17.8327
10.1331 | -34.529 | -17.8424
10.1392 | -34.5294 | -17.852
10.1453 | -34.5299 | -17.8618
10.1514 | -34.5303 | -17.8716
10.1575 | -34.5308 | -17.8815
10.1636 | -34.5312 | -17.891
10.1697 | -34.5317 | -17.9009
10.1758 | -34.5321 | -17.9104
10.1819 | -34.5326 | -17.9202
10.188 | -34.533 | -17.93
10.1941 | -34.5335 | -17.9397
10.2002 | -34.5339 | -17.9495
10.2063 | -34.5344 | -17.9592
10.2124 | -34.5348 | -17.9691
10.2185 | -34.5353 | -17.9789
10.2246 | -34.5357 | -17.9885
10.2307 | -34.5362 | -17.9984
10.2368 | -34.5366 | -18.0078
10.2429 | -34.5371 | -18.0176
10.2491 | -34.5375 | -18.0274
10.2552 | -34.5379 | -18.0372
10.2613 | -34.5384 | -18.0469
10.2674 | -34.5388 | -18.0567
10.2735 | -34.5393 | -18.0666
10.2796 | -34.5398 | -18.0763
10.2857 | -34.5402 | -18.0862
10.2918 | -34.5407 | -18.0961
10.2979 | -34.5411 | -18.1058
10.304 | -34.5416 | -18.1153
10.3101 | -34.542 | -18.1251
10.3162 | -34.5425 | -18.135
10.3223 | -34.5429 | -18.1445
10.3284 | -34.5434 | -18.1542
10.3345 | -34.5438 | -18.1638
10.3406 | -34.5443 | -18.1735
10.3467 | -34.5447 | -18.1832
10.3528 | -34.5451 | -18.1929
10.3589 | -34.5456 | -18.2026
10.365 | -34.546 | -18.2124
10.3711 | -34.5465 | -18.2222
10.3772 | -34.5469 | -18.2318
10.3833 | -34.5474 | -18.2415
10.3895 | -34.5479 | -18.2512
10.3956 | -34.5483 | -18.261
10.4017 | -34.5488 | -18.2705
10.4078 | -34.5492 | -18.2802
10.4139 | -34.5497 | -18.29
10.42 | -34.5501 | -18.2999
10.4261 | -34.5506 | -18.3095
10.4322 | -34.551 | -18.3192
10.4383 | -34.5514 | -18.3289
10.4444 | -34.5519 | -18.3384
10.4505 | -34.5523 | -18.3482
10.4566 | -34.5528 | -18.3578
10.4627 | -34.5532 | -18.3674
10.4688 | -34.5537 | -18.377
10.4749 | -34.5541 | -18.3868
10.481 | -34.5546 | -18.3965
10.4871 | -34.555 | -18.4061
10.4932 | -34.5555 | -18.4157
10.4993 | -34.5559 | -18.4254
10.5054 | -34.5564 | -18.4351
10.5115 | -34.5568 | -18.4448
10.5176 | -34.5573 | -18.4543
10.5237 | -34.5578 | -18.4642
10.5298 | -34.5582 | -18.4738
10.536 | -34.5586 | -18.4835
10.5421 | -34.5591 | -18.4931
10.5482 | -34.5595 | -18.5029
10.5543 | -34.56 | -18.5125
10.5604 | -34.5604 | -18.522
10.5665 | -34.5609 | -18.5316
10.5726 | -34.5613 | -18.5413
10.5787 | -34.5618 | -18.5509
10.5848 | -34.5622 | -18.5603
10.5909 | -34.5627 | -18.5701
10.597 | -34.5631 | -18.5798
10.6031 | -34.5636 | -18.5898
10.6092 | -34.564 | -18.5994
10.6153 | -34.5645 | -18.6092
10.6214 | -34.5649 | -18.6187
10.6275 | -34.5654 | -18.6284
10.6336 | -34.5658 | -18.6381
10.6397 | -34.5663 | -18.6477
10.6458 | -34.5667 | -18.6573
10.6519 | -34.5672 | -18.6671
10.658 | -34.5676 | -18.6767
10.6641 | -34.5681 | -18.6865
10.6702 | -34.5685 | -18.6962
10.6764 | -34.569 | -18.7059
10.6825 | -34.5694 | -18.7155
10.6886 | -34.5699 | -18.725
10.6947 | -34.5703 | -18.7346
10.7008 | -34.5708 | -18.7443
10.7069 | -34.5712 | -18.754
10.713 | -34.5717 | -18.7636
10.7191 | -34.5721 | -18.7733
10.7252 | -34.5726 | -18.7826
10.7313 | -34.573 | -18.7923
10.7374 | -34.5735 | -18.8019
10.7435 | -34.5739 | -18.8116
10.7496 | -34.5743 | -18.8211
10.7557 | -34.5748 | -18.8309
10.7618 | -34.5752 | -18.8403
10.7679 | -34.5757 | -18.8499
10.774 | -34.5761 | -18.8596
10.7801 | -34.5766 | -18.8693
10.7862 | -34.577 | -18.8787
10.7923 | -34.5775 | -18.8885
10.7984 | -34.5779 | -18.8979
10.8045 | -34.5784 | -18.9074
10.8106 | -34.5788 | -18.9171
10.8168 | -34.5793 | -18.9267
10.8229 | -34.5797 | -18.9364
10.829 | -34.5802 | -18.946
10.8351 | -34.5806 | -18.9554
10.8412 | -34.5811 | -18.9649
10.8473 | -34.5816 | -18.9743
10.8534 | -34.582 | -18.9839
10.8595 | -34.5825 | -18.9935
10.8656 | -34.5829 | -19.003
10.8717 | -34.5834 | -19.0126
10.8778 | -34.5839 | -19.0222
10.8839 | -34.5843 | -19.0319
10.89 | -34.5848 | -19.0416
10.8961 | -34.5852 | -19.0511
10.9022 | -34.5857 | -19.0607
10.9083 | -34.5861 | -19.0702
10.9144 | -34.5866 | -19.0797
10.9205 | -34.587 | -19.0893
10.9266 | -34.5875 | -19.0988
10.9327 | -34.5879 | -19.1083
10.9388 | -34.5884 | -19.1178
10.9449 | -34.5889 | -19.1273
10.951 | -34.5893 | -19.1368
10.9571 | -34.5898 | -19.1464
10.9633 | -34.5902 | -19.1561
10.9694 | -34.5906 | -19.1658
10.9755 | -34.5911 | -19.1754
10.9816 | -34.5916 | -19.1851
10.9877 | -34.592 | -19.1945
10.9938 | -34.5925 | -19.2042
10.9999 | -34.593 | -19.2136
11.006 | -34.5934 | -19.223
11.0121 | -34.5939 | -19.2326
11.0182 | -34.5943 | -19.2421
11.0243 | -34.5948 | -19.2517
11.0304 | -34.5952 | -19.2612
11.0365 | -34.5957 | -19.2707
11.0426 | -34.5962 | -19.2803
11.0487 | -34.5966 | -19.29
11.0548 | -34.5971 | -19.2993
11.0609 | -34.5976 | -19.309
11.067 | -34.598 | -19.3186
11.0731 | -34.5985 | -19.328
11.0792 | -34.599 | -19.3375
11.0853 | -34.5994 | -19.347
11.0914 | -34.5999 | -19.3566
11.0975 | -34.6003 | -19.3662
11.1037 | -34.6008 | -19.3757
11.1098 | -34.6012 | -19.3853
11.1159 | -34.6017 | -19.3946
11.122 | -34.6021 | -19.4041
11.1281 | -34.6026 | -19.4136
11.1342 | -34.6031 | -19.4231
11.1403 | -34.6035 | -19.4328
11.1464 | -34.604 | -19.4424
11.1525 | -34.6045 | -19.4517
11.1586 | -34.6049 | -19.4613
11.1647 | -34.6054 | -19.4707
11.1708 | -34.6058 | -19.4803
11.1769 | -34.6063 | -19.4898
11.183 | -34.6067 | -19.4991
11.1891 | -34.6072 | -19.5086
11.1952 | -34.6077 | -19.5181
11.2013 | -34.6081 | -19.5276
11.2074 | -34.6086 | -19.5371
11.2135 | -34.6091 | -19.5468
11.2196 | -34.6095 | -19.5563
11.2257 | -34.61 | -19.5659
11.2318 | -34.6105 | -19.5753
11.2379 | -34.611 | -19.585
11.244 | -34.6114 | -19.5946
11.2502 | -34.6119 | -19.604
11.2563 | -34.6124 | -19.6135
11.2624 | -34.6129 | -19.6229
11.2685 | -34.6133 | -19.6323
11.2746 | -34.6138 | -19.6418
11.2807 | -34.6143 | -19.6514
11.2868 | -34.6147 | -19.661
11.2929 | -34.6152 | -19.6702
11.299 | -34.6157 | -19.6797
11.3051 | -34.6161 | -19.6893
11.3112 | -34.6166 | -19.6989
11.3173 | -34.6171 | -19.7082
11.3234 | -34.6176 | -19.7175
11.3295 | -34.6181 | -19.7271
11.3356 | -34.6185 | -19.7366
11.3417 | -34.619 | -19.746
11.3478 | -34.6194 | -19.7557
11.3539 | -34.6199 | -19.7652
11.36 | -34.6203 | -19.7745
11.3661 | -34.6208 | -19.7839
11.3722 | -34.6213 | -19.7932
11.3783 | -34.6218 | -19.8027
11.3844 | -34.6222 | -19.8122
11.3906 | -34.6227 | -19.8217
11.3967 | -34.6232 | -19.831
11.4028 | -34.6237 | -19.8406
11.4089 | -34.6241 | -19.8499
11.415 | -34.6246 | -19.8594
11.4211 | -34.6251 | -19.8688
11.4272 | -34.6255 | -19.8783
11.4333 | -34.626 | -19.8875
11.4394 | -34.6264 | -19.8969
11.4455 | -34.6269 | -19.9063
11.4516 | -34.6274 | -19.9158
11.4577 | -34.6278 | -19.9252
11.4638 | -34.6283 | -19.9345
11.4699 | -34.6288 | -19.944
11.476 | -34.6293 | -19.9537
11.4821 | -34.6298 | -19.9631
11.4882 | -34.6302 | -19.9727
11.4943 | -34.6307 | -19.9822
11.5004 | -34.6312 | -19.9914
11.5065 | -34.6317 | -20.0008
11.5126 | -34.6322 | -20.0101
11.5187 | -34.6326 | -20.0196
11.5248 | -34.6331 | -20.0292
11.5309 | -34.6336 | -20.0386
11.5371 | -34.6341 | -20.0479
11.5432 | -34.6345 | -20.0574
11.5493 | -34.635 | -20.067
11.5554 | -34.6355 | -20.0764
11.5615 | -34.636 | -20.086
11.5676 | -34.6365 | -20.0954
11.5737 | -34.6369 | -20.1047
11.5798 | -34.6374 | -20.114
11.5859 | -34.6379 | -20.1234
11.592 | -34.6384 | -20.1328
11.5981 | -34.6389 | -20.1422
11.6042 | -34.6394 | -20.1516
11.6103 | -34.6398 | -20.1609
11.6164 | -34.6403 | -20.1702
11.6225 | -34.6408 | -20.1796
11.6286 | -34.6412 | -20.189
11.6347 | -34.6417 | -20.1984
11.6408 | -34.6422 | -20.2077
11.6469 | -34.6427 | -20.2171
11.653 | -34.6431 | -20.2264
11.6591 | -34.6436 | -20.2359
11.6652 | -34.6441 | -20.2452
11.6713 | -34.6446 | -20.2545
11.6775 | -34.6451 | -20.2638
11.6836 | -34.6456 | -20.2731
11.6897 | -34.6461 | -20.2825
11.6958 | -34.6465 | -20.292
11.7019 | -34.647 | -20.3015
11.708 | -34.6475 | -20.3108
11.7141 | -34.648 | -20.32
11.7202 | -34.6485 | -20.3294
11.7263 | -34.6489 | -20.3389
11.7324 | -34.6494 | -20.3484
11.7385 | -34.6499 | -20.3578
11.7446 | -34.6503 | -20.3672
11.7507 | -34.6509 | -20.3764
11.7568 | -34.6513 | -20.3858
11.7629 | -34.6518 | -20.3952
11.769 | -34.6523 | -20.4045
11.7751 | -34.6528 | -20.4138
11.7812 | -34.6533 | -20.423
11.7873 | -34.6538 | -20.4324
11.7934 | -34.6542 | -20.4418
11.7995 | -34.6547 | -20.4511
11.8056 | -34.6552 | -20.4604
11.8117 | -34.6557 | -20.4698
11.8178 | -34.6562 | -20.479
11.824 | -34.6567 | -20.4882
11.8301 | -34.6572 | -20.4975
11.8362 | -34.6576 | -20.5067
11.8423 | -34.6581 | -20.5161
11.8484 | -34.6586 | -20.5255
11.8545 | -34.6591 | -20.5346
11.8606 | -34.6596 | -20.5439
11.8667 | -34.6601 | -20.5533
11.8728 | -34.6606 | -20.5624
11.8789 | -34.661 | -20.5717
11.885 | -34.6615 | -20.5811
11.8911 | -34.662 | -20.5903
11.8972 | -34.6625 | -20.5996
11.9033 | -34.663 | -20.6088
11.9094 | -34.6635 | -20.6181
11.9155 | -34.664 | -20.6273
11.9216 | -34.6645 | -20.6365
11.9277 | -34.665 | -20.6457
11.9338 | -34.6654 | -20.655
11.9399 | -34.6659 | -20.6644
11.946 | -34.6664 | -20.6736
11.9521 | -34.6669 | -20.6828
11.9582 | -34.6674 | -20.692
11.9644 | -34.6679 | -20.7012
11.9705 | -34.6684 | -20.7104
11.9766 | -34.6689 | -20.7197
11.9827 | -34.6694 | -20.7289
11.9888 | -34.6699 | -20.7381
11.9949 | -34.6704 | -20.7471
12.001 | -34.6709 | -20.7564
12.0071 | -34.6714 | -20.7656
12.0132 | -34.6719 | -20.7748
12.0193 | -34.6723 | -20.7841
12.0254 | -34.6728 | -20.7933
12.0315 | -34.6733 | -20.8023
12.0376 | -34.6738 | -20.8115
12.0437 | -34.6743 | -20.8207
12.0498 | -34.6748 | -20.83
12.0559 | -34.6752 | -20.8391
12.062 | -34.6758 | -20.8481
12.0681 | -34.6763 | -20.8574
12.0742 | -34.6768 | -20.8665
12.0803 | -34.6772 | -20.8759
12.0864 | -34.6777 | -20.8852
12.0925 | -34.6782 | -20.8944
12.0986 | -34.6788 | -20.9035
12.1047 | -34.6792 | -20.9128
12.1109 | -34.6798 | -20.9219
12.117 | -34.6802 | -20.931
12.1231 | -34.6808 | -20.9402
12.1292 | -34.6812 | -20.9494
12.1353 | -34.6817 | -20.9585
12.1414 | -34.6822 | -20.9677
12.1475 | -34.6827 | -20.977
12.1536 | -34.6832 | -20.9863
12.1597 | -34.6837 | -20.9954
12.1658 | -34.6842 | -21.0045
12.1719 | -34.6847 | -21.0135
12.178 | -34.6852 | -21.0227
12.1841 | -34.6857 | -21.0319
12.1902 | -34.6862 | -21.0411
12.1963 | -34.6867 | -21.0504
12.2024 | -34.6872 | -21.0593
12.2085 | -34.6877 | -21.0685
12.2146 | -34.6882 | -21.0777
12.2207 | -34.6887 | -21.0868
12.2268 | -34.6892 | -21.096
12.2329 | -34.6897 | -21.105
12.239 | -34.6902 | -21.1138
12.2451 | -34.6907 | -21.123
12.2513 | -34.6912 | -21.1323
12.2574 | -34.6917 | -21.1414
12.2635 | -34.6922 | -21.1506
12.2696 | -34.6927 | -21.1596
12.2757 | -34.6932 | -21.1685
12.2818 | -34.6937 | -21.1777
12.2879 | -34.6942 | -21.1869
12.294 | -34.6947 | -21.1959
12.3001 | -34.6952 | -21.2051
12.3062 | -34.6957 | -21.2142
12.3123 | -34.6962 | -21.223
12.3184 | -34.6967 | -21.2321
12.3245 | -34.6972 | -21.241
12.3306 | -34.6977 | -21.2502
12.3367 | -34.6982 | -21.2592
12.3428 | -34.6987 | -21.2683
12.3489 | -34.6993 | -21.2771
12.355 | -34.6997 | -21.2864
12.3611 | -34.7003 | -21.2956
12.3672 | -34.7007 | -21.3047
12.3733 | -34.7013 | -21.3138
12.3794 | -34.7018 | -21.3226
12.3855 | -34.7022 | -21.3316
12.3916 | -34.7027 | -21.3407
12.3978 | -34.7033 | -21.3499
12.4039 | -34.7038 | -21.359
12.41 | -34.7043 | -21.3681
12.4161 | -34.7048 | -21.3768
12.4222 | -34.7053 | -21.3858
12.4283 | -34.7058 | -21.3948
12.4344 | -34.7063 | -21.404
12.4405 | -34.7068 | -21.4131
12.4466 | -34.7073 | -21.4221
12.4527 | -34.7078 | -21.431
12.4588 | -34.7083 | -21.4401
12.4649 | -34.7088 | -21.4491
12.471 | -34.7093 | -21.4583
12.4771 | -34.7098 | -21.4673
12.4832 | -34.7103 | -21.4764
12.4893 | -34.7108 | -21.4852
12.4954 | -34.7113 | -21.4941
12.5015 | -34.7118 | -21.5031
12.5076 | -34.7123 | -21.5122
12.5137 | -34.7128 | -21.5213
12.5198 | -34.7133 | -21.5302
12.5259 | -34.7138 | -21.5392
12.532 | -34.7143 | -21.5482
12.5382 | -34.7148 | -21.5573
12.5443 | -34.7153 | -21.5663
12.5504 | -34.7158 | -21.5754
12.5565 | -34.7163 | -21.5842
12.5626 | -34.7169 | -21.5933
12.5687 | -34.7173 | -21.6024
12.5748 | -34.7178 | -21.6114
12.5809 | -34.7183 | -21.6204
12.587 | -34.7188 | -21.6295
12.5931 | -34.7194 | -21.6383
12.5992 | -34.7199 | -21.6473
12.6053 | -34.7204 | -21.6563
12.6114 | -34.7209 | -21.6652
12.6175 | -34.7214 | -21.6744
12.6236 | -34.7219 | -21.6834
12.6297 | -34.7224 | -21.6923
12.6358 | -34.7229 | -21.7011
12.6419 | -34.7234 | -21.7102
12.648 | -34.724 | -21.7192
12.6541 | -34.7244 | -21.7282
12.6602 | -34.7249 | -21.7367
12.6663 | -34.7254 | -21.7457
12.6724 | -34.726 | -21.7546
12.6785 | -34.7265 | -21.7635
12.6847 | -34.727 | -21.7724
12.6908 | -34.7275 | -21.7813
12.6969 | -34.728 | -21.7901
12.703 | -34.7285 | -21.7992
12.7091 | -34.729 | -21.8082
12.7152 | -34.7295 | -21.817
12.7213 | -34.73 | -21.826
12.7274 | -34.7305 | -21.835
12.7335 | -34.731 | -21.8436
12.7396 | -34.7315 | -21.8525
12.7457 | -34.732 | -21.8615
12.7518 | -34.7325 | -21.8704
12.7579 | -34.733 | -21.8794
12.764 | -34.7335 | -21.8885
12.7701 | -34.734 | -21.8974
12.7762 | -34.7345 | -21.9063
12.7823 | -34.7351 | -21.9152
12.7884 | -34.7356 | -21.9241
12.7945 | -34.7361 | -21.9333
12.8006 | -34.7366 | -21.9421
12.8067 | -34.7371 | -21.9509
12.8128 | -34.7376 | -21.9597
12.8189 | -34.7381 | -21.9686
12.8251 | -34.7386 | -21.9776
12.8312 | -34.7391 | -21.9866
12.8373 | -34.7396 | -21.9954
12.8434 | -34.7401 | -22.0044
12.8495 | -34.7406 | -22.0132
12.8556 | -34.7411 | -22.022
12.8617 | -34.7416 | -22.0309
12.8678 | -34.7421 | -22.04
12.8739 | -34.7426 | -22.0487
12.88 | -34.7431 | -22.0576
12.8861 | -34.7436 | -22.0664
12.8922 | -34.7441 | -22.0753
12.8983 | -34.7446 | -22.0842
12.9044 | -34.7451 | -22.0929
12.9105 | -34.7456 | -22.1017
12.9166 | -34.7461 | -22.1106
12.9227 | -34.7466 | -22.1194
12.9288 | -34.7471 | -22.1283
12.9349 | -34.7476 | -22.1372
12.941 | -34.7481 | -22.1461
12.9471 | -34.7486 | -22.155
12.9532 | -34.7492 | -22.1639
12.9593 | -34.7497 | -22.1729
12.9654 | -34.7502 | -22.1819
12.9716 | -34.7507 | -22.1908
12.9777 | -34.7512 | -22.1997
12.9838 | -34.7517 | -22.2085
12.9899 | -34.7522 | -22.2174
12.996 | -34.7527 | -22.2262
13.0021 | -34.7533 | -22.2351
13.0082 | -34.7538 | -22.2439
13.0143 | -34.7543 | -22.2525
13.0204 | -34.7548 | -22.2612
13.0265 | -34.7553 | -22.27
13.0326 | -34.7558 | -22.2789
13.0387 | -34.7563 | -22.2879
13.0448 | -34.7568 | -22.2968
13.0509 | -34.7573 | -22.3055
13.057 | -34.7578 | -22.3143
13.0631 | -34.7584 | -22.3231
13.0692 | -34.7589 | -22.3319
13.0753 | -34.7594 | -22.3409
13.0814 | -34.7599 | -22.3498
13.0875 | -34.7604 | -22.3585
13.0936 | -34.7609 | -22.3674
13.0997 | -34.7614 | -22.3762
13.1058 | -34.7619 | -22.3849
13.112 | -34.7624 | -22.3937
13.1181 | -34.7629 | -22.4024
13.1242 | -34.7634 | -22.4113
13.1303 | -34.764 | -22.4202
13.1364 | -34.7645 | -22.429
13.1425 | -34.7649 | -22.4378
13.1486 | -34.7654 | -22.4466
13.1547 | -34.766 | -22.4554
13.1608 | -34.7665 | -22.4642
13.1669 | -34.767 | -22.4729
13.173 | -34.7675 | -22.4818
13.1791 | -34.768 | -22.4905
13.1852 | -34.7685 | -22.4993
13.1913 | -34.769 | -22.5079
13.1974 | -34.7695 | -22.5168
13.2035 | -34.77 | -22.5256
13.2096 | -34.7705 | -22.5343
13.2157 | -34.771 | -22.5431
13.2218 | -34.7715 | -22.5518
13.2279 | -34.7721 | -22.5604
13.234 | -34.7726 | -22.5692
13.2401 | -34.7731 | -22.5779
13.2462 | -34.7736 | -22.5867
13.2524 | -34.7741 | -22.5955
13.2585 | -34.7746 | -22.6042
13.2646 | -34.7751 | -22.613
13.2707 | -34.7756 | -22.6219
13.2768 | -34.7761 | -22.6308
13.2829 | -34.7766 | -22.6397
13.289 | -34.7772 | -22.6483
13.2951 | -34.7777 | -22.657
13.3012 | -34.7782 | -22.6657
13.3073 | -34.7787 | -22.6746
13.3134 | -34.7792 | -22.6834
13.3195 | -34.7797 | -22.6921
13.3256 | -34.7802 | -22.7008
13.3317 | -34.7807 | -22.7095
13.3378 | -34.7812 | -22.7184
13.3439 | -34.7817 | -22.7273
13.35 | -34.7822 | -22.7361
13.3561 | -34.7827 | -22.7448
13.3622 | -34.7832 | -22.7535
13.3683 | -34.7837 | -22.7621
13.3744 | -34.7842 | -22.7708
13.3805 | -34.7847 | -22.7796
13.3866 | -34.7852 | -22.7882
13.3927 | -34.7857 | -22.797
13.3989 | -34.7862 | -22.8057
13.405 | -34.7867 | -22.8145
13.4111 | -34.7872 | -22.8232
13.4172 | -34.7877 | -22.832
13.4233 | -34.7882 | -22.8408
13.4294 | -34.7887 | -22.8498
13.4355 | -34.7892 | -22.8583
13.4416 | -34.7897 | -22.867
13.4477 | -34.7903 | -22.8757
13.4538 | -34.7908 | -22.8845
13.4599 | -34.7913 | -22.8933
13.466 | -34.7918 | -22.9022
13.4721 | -34.7923 | -22.9108
13.4782 | -34.7928 | -22.9195
13.4843 | -34.7933 | -22.9284
13.4904 | -34.7938 | -22.937
13.4965 | -34.7943 | -22.9459
13.5026 | -34.7948 | -22.9547
13.5087 | -34.7953 | -22.9633
13.5148 | -34.7958 | -22.9722
13.5209 | -34.7963 | -22.9811
13.527 | -34.7968 | -22.9898
13.5331 | -34.7974 | -22.9985
13.5393 | -34.7979 | -23.0071
13.5454 | -34.7984 | -23.0158
13.5515 | -34.7989 | -23.0247
13.5576 | -34.7994 | -23.0334
13.5637 | -34.7999 | -23.0423
13.5698 | -34.8004 | -23.051
13.5759 | -34.8009 | -23.0594
13.582 | -34.8014 | -23.0681
13.5881 | -34.8019 | -23.0768
13.5942 | -34.8024 | -23.0857
13.6003 | -34.803 | -23.0944
13.6064 | -34.8035 | -23.1032
13.6125 | -34.804 | -23.112
13.6186 | -34.8045 | -23.1207
13.6247 | -34.805 | -23.1295
13.6308 | -34.8055 | -23.1382
13.6369 | -34.806 | -23.1471
13.643 | -34.8065 | -23.1559
13.6491 | -34.807 | -23.1645
13.6552 | -34.8075 | -23.1732
13.6613 | -34.808 | -23.1819
13.6674 | -34.8085 | -23.1906
13.6735 | -34.809 | -23.1994
13.6796 | -34.8096 | -23.2081
13.6858 | -34.81 | -23.2167
13.6919 | -34.8106 | -23.2254
13.698 | -34.8111 | -23.2341
13.7041 | -34.8116 | -23.243
13.7102 | -34.8121 | -23.2517
13.7163 | -34.8126 | -23.2603
13.7224 | -34.8131 | -23.2691
13.7285 | -34.8136 | -23.2778
13.7346 | -34.8141 | -23.2865
13.7407 | -34.8147 | -23.2952
13.7468 | -34.8152 | -23.304
13.7529 | -34.8157 | -23.3126
13.759 | -34.8162 | -23.3214
13.7651 | -34.8167 | -23.3301
13.7712 | -34.8172 | -23.3389
13.7773 | -34.8177 | -23.3476
13.7834 | -34.8182 | -23.3563
13.7895 | -34.8187 | -23.365
13.7956 | -34.8192 | -23.3738
13.8017 | -34.8197 | -23.3824
13.8078 | -34.8202 | -23.3911
13.8139 | -34.8207 | -23.3997
13.82 | -34.8212 | -23.4084
13.8262 | -34.8217 | -23.4171
13.8323 | -34.8223 | -23.4259
13.8384 | -34.8228 | -23.4346
13.8445 | -34.8233 | -23.4433
13.8506 | -34.8238 | -23.452
13.8567 | -34.8243 | -23.4608
13.8628 | -34.8248 | -23.4695
13.8689 | -34.8253 | -23.4783
13.875 | -34.8258 | -23.4869
13.8811 | -34.8263 | -23.4957
13.8872 | -34.8268 | -23.5043
13.8933 | -34.8273 | -23.5129
13.8994 | -34.8278 | -23.5216
13.9055 | -34.8283 | -23.5305
13.9116 | -34.8288 | -23.5391
13.9177 | -34.8293 | -23.5478
13.9238 | -34.8298 | -23.5565
13.9299 | -34.8303 | -23.5651
13.936 | -34.8308 | -23.5739
13.9421 | -34.8313 | -23.5825
13.9482 | -34.8318 | -23.5911
13.9543 | -34.8323 | -23.5999
13.9604 | -34.8328 | -23.6086
13.9665 | -34.8334 | -23.6171
13.9727 | -34.8339 | -23.626
13.9788 | -34.8344 | -23.6347
13.9849 | -34.8349 | -23.6434
13.991 | -34.8354 | -23.6521
13.9971 | -34.8359 | -23.6607
14.0032 | -34.8364 | -23.6694
14.0093 | -34.8369 | -23.6781
14.0154 | -34.8374 | -23.6867
14.0215 | -34.8379 | -23.6953
14.0276 | -34.8385 | -23.704
14.0337 | -34.839 | -23.7127
14.0398 | -34.8395 | -23.7213
14.0459 | -34.84 | -23.73
14.052 | -34.8405 | -23.7388
14.0581 | -34.841 | -23.7475
14.0642 | -34.8415 | -23.7562
14.0703 | -34.842 | -23.7649
14.0764 | -34.8425 | -23.7735
14.0825 | -34.843 | -23.7822
14.0886 | -34.8436 | -23.7909
14.0947 | -34.8441 | -23.7996
14.1008 | -34.8446 | -23.8082
14.1069 | -34.8451 | -23.8167
14.1131 | -34.8456 | -23.8255
14.1192 | -34.8461 | -23.8341
14.1253 | -34.8466 | -23.8426
14.1314 | -34.8471 | -23.8513
14.1375 | -34.8476 | -23.8598
14.1436 | -34.8482 | -23.8685
14.1497 | -34.8487 | -23.8772
14.1558 | -34.8492 | -23.8859
14.1619 | -34.8497 | -23.8945
14.168 | -34.8502 | -23.9031
14.1741 | -34.8507 | -23.9117
14.1802 | -34.8512 | -23.9203
14.1863 | -34.8517 | -23.929
14.1924 | -34.8523 | -23.9377
14.1985 | -34.8528 | -23.9463
14.2046 | -34.8533 | -23.9551
14.2107 | -34.8538 | -23.9635
14.2168 | -34.8543 | -23.9722
14.2229 | -34.8548 | -23.9809
14.229 | -34.8553 | -23.9896
14.2351 | -34.8558 | -23.9982
14.2412 | -34.8563 | -24.0069
14.2473 | -34.8569 | -24.0155
14.2534 | -34.8574 | -24.024
14.2596 | -34.8579 | -24.0325
14.2657 | -34.8584 | -24.0412
14.2718 | -34.8589 | -24.0499
14.2779 | -34.8594 | -24.0585
14.284 | -34.8599 | -24.0672
14.2901 | -34.8604 | -24.0758
14.2962 | -34.861 | -24.0845
14.3023 | -34.8614 | -24.0932
14.3084 | -34.862 | -24.1018
14.3145 | -34.8625 | -24.1104
14.3206 | -34.863 | -24.1192
14.3267 | -34.8635 | -24.1278
14.3328 | -34.864 | -24.1364
14.3389 | -34.8646 | -24.145
14.345 | -34.865 | -24.1536
14.3511 | -34.8656 | -24.1622
14.3572 | -34.8661 | -24.171
14.3633 | -34.8666 | -24.1797
14.3694 | -34.8671 | -24.1883
14.3755 | -34.8676 | -24.1969
14.3816 | -34.8682 | -24.2057
14.3877 | -34.8687 | -24.2142
14.3938 | -34.8692 | -24.223
14.4 | -34.8697 | -24.2315
14.4061 | -34.8702 | -24.2401
14.4122 | -34.8708 | -24.2489
14.4183 | -34.8713 | -24.2574
14.4244 | -34.8718 | -24.266
14.4305 | -34.8723 | -24.2746
14.4366 | -34.8729 | -24.2833
14.4427 | -34.8734 | -24.2921
14.4488 | -34.8739 | -24.3007
14.4549 | -34.8744 | -24.3093
14.461 | -34.8749 | -24.3178
14.4671 | -34.8754 | -24.3265
14.4732 | -34.876 | -24.3353
14.4793 | -34.8765 | -24.3439
14.4854 | -34.877 | -24.3524
14.4915 | -34.8775 | -24.3608
14.4976 | -34.8781 | -24.3693
14.5037 | -34.8786 | -24.3779
14.5098 | -34.8791 | -24.3865
14.5159 | -34.8796 | -24.3951
14.522 | -34.8801 | -24.4037
14.5281 | -34.8806 | -24.4122
14.5342 | -34.8812 | -24.4209
14.5403 | -34.8817 | -24.4294
14.5465 | -34.8822 | -24.4381
14.5526 | -34.8827 | -24.4468
14.5587 | -34.8832 | -24.4555
14.5648 | -34.8838 | -24.464
14.5709 | -34.8843 | -24.4725
14.577 | -34.8848 | -24.481
14.5831 | -34.8853 | -24.4896
14.5892 | -34.8858 | -24.4982
14.5953 | -34.8864 | -24.5067
14.6014 | -34.8869 | -24.5153
14.6075 | -34.8874 | -24.5238
14.6136 | -34.8879 | -24.5325
14.6197 | -34.8884 | -24.5412
14.6258 | -34.8889 | -24.5498
14.6319 | -34.8894 | -24.5583
14.638 | -34.8899 | -24.5669
14.6441 | -34.8905 | -24.5755
14.6502 | -34.891 | -24.5843
14.6563 | -34.8915 | -24.5929
14.6624 | -34.892 | -24.6015
14.6685 | -34.8926 | -24.6099
14.6746 | -34.8931 | -24.6185
14.6807 | -34.8936 | -24.627
14.6869 | -34.8941 | -24.6356
14.693 | -34.8947 | -24.6441
14.6991 | -34.8952 | -24.6527
14.7052 | -34.8957 | -24.6611
14.7113 | -34.8963 | -24.6697
14.7174 | -34.8968 | -24.6784
14.7235 | -34.8973 | -24.6869
14.7296 | -34.8978 | -24.6954
14.7357 | -34.8984 | -24.7039
14.7418 | -34.8989 | -24.7124
14.7479 | -34.8994 | -24.7209
14.754 | -34.8999 | -24.7296
14.7601 | -34.9005 | -24.7381
14.7662 | -34.901 | -24.7466
14.7723 | -34.9015 | -24.7549
14.7784 | -34.902 | -24.7636
14.7845 | -34.9025 | -24.7721
14.7906 | -34.9031 | -24.7808
14.7967 | -34.9036 | -24.7893
14.8028 | -34.9041 | -24.7978
14.8089 | -34.9046 | -24.8063
14.815 | -34.9052 | -24.8149
14.8211 | -34.9057 | -24.8234
14.8272 | -34.9062 | -24.832
14.8334 | -34.9067 | -24.8405
14.8395 | -34.9073 | -24.849
14.8456 | -34.9078 | -24.8574
14.8517 | -34.9083 | -24.8661
14.8578 | -34.9088 | -24.8747
14.8639 | -34.9094 | -24.8833
14.87 | -34.9099 | -24.8918
14.8761 | -34.9104 | -24.9003
14.8822 | -34.911 | -24.9088
14.8883 | -34.9115 | -24.9174
14.8944 | -34.912 | -24.9259
14.9005 | -34.9125 | -24.9344
14.9066 | -34.913 | -24.943
14.9127 | -34.9136 | -24.9514
14.9188 | -34.9141 | -24.9599
14.9249 | -34.9146 | -24.9685
14.931 | -34.9152 | -24.9771
14.9371 | -34.9157 | -24.9855
14.9432 | -34.9162 | -24.9941
14.9493 | -34.9168 | -25.0025
14.9554 | -34.9173 | -25.011
14.9615 | -34.9178 | -25.0195
14.9676 | -34.9183 | -25.0282
14.9738 | -34.9188 | -25.0366
14.9799 | -34.9194 | -25.0452
14.986 | -34.9199 | -25.0536
14.9921 | -34.9204 | -25.0621
14.9982 | -34.9209 | -25.0705
15.0043 | -34.9215 | -25.079
15.0104 | -34.922 | -25.0876
15.0165 | -34.9225 | -25.0962
15.0226 | -34.9231 | -25.1046
15.0287 | -34.9236 | -25.1131
15.0348 | -34.9241 | -25.1217
15.0409 | -34.9246 | -25.1302
15.047 | -34.9252 | -25.1388
15.0531 | -34.9257 | -25.1472
15.0592 | -34.9263 | -25.1557
15.0653 | -34.9268 | -25.1642
15.0714 | -34.9273 | -25.1726
15.0775 | -34.9278 | -25.181
15.0836 | -34.9284 | -25.1895
15.0897 | -34.929 | -25.1978
15.0958 | -34.9295 | -25.2062
15.1019 | -34.93 | -25.2147
15.108 | -34.9305 | -25.2231
15.1141 | -34.9311 | -25.2317
15.1203 | -34.9316 | -25.2402
15.1264 | -34.9322 | -25.2486
15.1325 | -34.9327 | -25.2571
15.1386 | -34.9332 | -25.2655
15.1447 | -34.9338 | -25.274
15.1508 | -34.9343 | -25.2825
15.1569 | -34.9348 | -25.291
15.163 | -34.9354 | -25.2992
15.1691 | -34.9359 | -25.3076
15.1752 | -34.9365 | -25.3161
15.1813 | -34.937 | -25.3245
15.1874 | -34.9375 | -25.333
15.1935 | -34.9381 | -25.3413
15.1996 | -34.9386 | -25.3497
15.2057 | -34.9391 | -25.3582
15.2118 | -34.9397 | -25.3667
15.2179 | -34.9402 | -25.3752
15.224 | -34.9408 | -25.3836
15.2301 | -34.9413 | -25.3919
15.2362 | -34.9419 | -25.4004
15.2423 | -34.9424 | -25.4089
15.2484 | -34.9429 | -25.4174
15.2545 | -34.9435 | -25.4259
15.2607 | -34.944 | -25.4344
15.2668 | -34.9446 | -25.4426
15.2729 | -34.9451 | -25.451
15.279 | -34.9457 | -25.4596
15.2851 | -34.9462 | -25.468
15.2912 | -34.9468 | -25.4765
15.2973 | -34.9473 | -25.485
15.3034 | -34.9478 | -25.4933
15.3095 | -34.9484 | -25.5018
15.3156 | -34.9489 | -25.5101
15.3217 | -34.9495 | -25.5185
15.3278 | -34.95 | -25.5271
15.3339 | -34.9506 | -25.5353
15.34 | -34.9511 | -25.5436
15.3461 | -34.9517 | -25.552
15.3522 | -34.9522 | -25.5604
15.3583 | -34.9527 | -25.569
15.3644 | -34.9533 | -25.5773
15.3705 | -34.9538 | -25.5854
15.3766 | -34.9544 | -25.5939
15.3827 | -34.9549 | -25.6022
15.3888 | -34.9555 | -25.6106
15.3949 | -34.956 | -25.619
15.401 | -34.9565 | -25.6275
15.4072 | -34.9571 | -25.6357
15.4133 | -34.9576 | -25.6442
15.4194 | -34.9582 | -25.6526
15.4255 | -34.9587 | -25.6611
15.4316 | -34.9592 | -25.6695
15.4377 | -34.9598 | -25.678
15.4438 | -34.9603 | -25.6863
15.4499 | -34.9609 | -25.6946
15.456 | -34.9614 | -25.7029
15.4621 | -34.962 | -25.7113
15.4682 | -34.9625 | -25.7197
15.4743 | -34.9631 | -25.7279
15.4804 | -34.9636 | -25.7363
15.4865 | -34.9641 | -25.7447
15.4926 | -34.9647 | -25.7532
15.4987 | -34.9652 | -25.7615
15.5048 | -34.9658 | -25.7699
15.5109 | -34.9663 | -25.7782
15.517 | -34.9669 | -25.7866
15.5231 | -34.9674 | -25.7951
15.5292 | -34.968 | -25.8034
15.5353 | -34.9685 | -25.8119
15.5414 | -34.969 | -25.8202
15.5476 | -34.9696 | -25.8283
15.5537 | -34.9702 | -25.8367
15.5598 | -34.9707 | -25.8452
15.5659 | -34.9712 | -25.8536
15.572 | -34.9718 | -25.8619
15.5781 | -34.9723 | -25.8702
15.5842 | -34.9729 | -25.8784
15.5903 | -34.9734 | -25.8866
15.5964 | -34.974 | -25.8949
15.6025 | -34.9746 | -25.9032
15.6086 | -34.9751 | -25.9116
15.6147 | -34.9756 | -25.92
15.6208 | -34.9762 | -25.9282
15.6269 | -34.9768 | -25.9364
15.633 | -34.9773 | -25.9449
15.6391 | -34.9779 | -25.9532
15.6452 | -34.9784 | -25.9616
15.6513 | -34.979 | -25.9697
15.6574 | -34.9795 | -25.9781
15.6635 | -34.9801 | -25.9864
15.6696 | -34.9806 | -25.9948
15.6757 | -34.9812 | -26.0032
15.6818 | -34.9817 | -26.0115
15.688 | -34.9823 | -26.0197
15.6941 | -34.9828 | -26.028
15.7002 | -34.9833 | -26.0363
15.7063 | -34.9839 | -26.0445
15.7124 | -34.9844 | -26.0529
15.7185 | -34.985 | -26.0612
15.7246 | -34.9855 | -26.0694
15.7307 | -34.9861 | -26.0777
15.7368 | -34.9866 | -26.086
15.7429 | -34.9872 | -26.0943
15.749 | -34.9878 | -26.1026
15.7551 | -34.9883 | -26.111
15.7612 | -34.9889 | -26.1192
15.7673 | -34.9894 | -26.1274
15.7734 | -34.99 | -26.1357
15.7795 | -34.9905 | -26.144
15.7856 | -34.9911 | -26.1522
15.7917 | -34.9917 | -26.1602
15.7978 | -34.9922 | -26.1686
15.8039 | -34.9928 | -26.1769
15.81 | -34.9933 | -26.1852
15.8161 | -34.9939 | -26.1935
15.8222 | -34.9944 | -26.2018
15.8283 | -34.995 | -26.2098
15.8345 | -34.9955 | -26.2181
15.8406 | -34.9961 | -26.2263
15.8467 | -34.9966 | -26.2346
15.8528 | -34.9972 | -26.243
15.8589 | -34.9978 | -26.2513
15.865 | -34.9984 | -26.2592
15.8711 | -34.9989 | -26.2676
15.8772 | -34.9995 | -26.2757
15.8833 | -35 | -26.2841
15.8894 | -35.0006 | -26.2924
15.8955 | -35.0011 | -26.3007
15.9016 | -35.0017 | -26.3088
15.9077 | -35.0022 | -26.3171
15.9138 | -35.0027 | -26.3254
15.9199 | -35.0033 | -26.3336
15.926 | -35.0038 | -26.3417
15.9321 | -35.0044 | -26.3498
15.9382 | -35.0049 | -26.358
15.9443 | -35.0055 | -26.3662
15.9504 | -35.006 | -26.3745
15.9565 | -35.0066 | -26.3827
15.9626 | -35.0071 | -26.391
15.9687 | -35.0077 | -26.399
15.9749 | -35.0082 | -26.4071
15.981 | -35.0088 | -26.4154
15.9871 | -35.0093 | -26.4237
15.9932 | -35.0099 | -26.4318
15.9993 | -35.0104 | -26.4401
16.0054 | -35.011 | -26.448
16.0115 | -35.0116 | -26.4562
16.0176 | -35.0121 | -26.4645
16.0237 | -35.0127 | -26.4728
16.0298 | -35.0132 | -26.481
16.0359 | -35.0138 | -26.4892
16.042 | -35.0143 | -26.4973
16.0481 | -35.0149 | -26.5055
16.0542 | -35.0155 | -26.5138
16.0603 | -35.016 | -26.5219
16.0664 | -35.0165 | -26.5302
16.0725 | -35.0171 | -26.5382
16.0786 | -35.0177 | -26.5463
16.0847 | -35.0182 | -26.5546
16.0908 | -35.0188 | -26.5628
16.0969 | -35.0193 | -26.5711
16.103 | -35.0199 | -26.5794
16.1091 | -35.0204 | -26.5874
16.1152 | -35.021 | -26.5956
16.1214 | -35.0215 | -26.6038
16.1275 | -35.0221 | -26.612
16.1336 | -35.0227 | -26.6201
16.1397 | -35.0232 | -26.6283
16.1458 | -35.0238 | -26.6363
16.1519 | -35.0244 | -26.6446
16.158 | -35.0249 | -26.6529
16.1641 | -35.0254 | -26.6611
16.1702 | -35.026 | -26.6693
16.1763 | -35.0266 | -26.6776
16.1824 | -35.0271 | -26.6855
16.1885 | -35.0277 | -26.6936
16.1946 | -35.0282 | -26.7017
16.2007 | -35.0288 | -26.7098
16.2068 | -35.0293 | -26.718
16.2129 | -35.0299 | -26.7262
16.219 | -35.0305 | -26.7341
16.2251 | -35.031 | -26.7422
16.2312 | -35.0316 | -26.7504
16.2373 | -35.0321 | -26.7584
16.2434 | -35.0327 | -26.7666
16.2495 | -35.0333 | -26.7746
16.2556 | -35.0339 | -26.7829
16.2618 | -35.0344 | -26.791
16.2679 | -35.035 | -26.7992
16.274 | -35.0355 | -26.8074
16.2801 | -35.0361 | -26.8155
16.2862 | -35.0366 | -26.8235
16.2923 | -35.0372 | -26.8317
16.2984 | -35.0377 | -26.8399
16.3045 | -35.0383 | -26.848
16.3106 | -35.0388 | -26.8561
16.3167 | -35.0394 | -26.8643
16.3228 | -35.04 | -26.8722
16.3289 | -35.0405 | -26.8803
16.335 | -35.041 | -26.8884
16.3411 | -35.0416 | -26.8965
16.3472 | -35.0422 | -26.9048
16.3533 | -35.0427 | -26.913
16.3594 | -35.0433 | -26.9208
16.3655 | -35.0438 | -26.9289
16.3716 | -35.0444 | -26.9369
16.3777 | -35.0449 | -26.945
16.3838 | -35.0455 | -26.9531
16.3899 | -35.0461 | -26.9611
16.396 | -35.0466 | -26.9692
16.4021 | -35.0472 | -26.9772
16.4083 | -35.0477 | -26.9852
16.4144 | -35.0483 | -26.9934
16.4205 | -35.0488 | -27.0015
16.4266 | -35.0494 | -27.0094
16.4327 | -35.05 | -27.0175
16.4388 | -35.0505 | -27.0256
16.4449 | -35.0511 | -27.0336
16.451 | -35.0516 | -27.0418
16.4571 | -35.0522 | -27.05
16.4632 | -35.0527 | -27.0578
16.4693 | -35.0533 | -27.0659
16.4754 | -35.0538 | -27.0741
16.4815 | -35.0544 | -27.0822
16.4876 | -35.0549 | -27.0902
16.4937 | -35.0555 | -27.0982
16.4998 | -35.056 | -27.106
16.5059 | -35.0566 | -27.114
16.512 | -35.0572 | -27.1223
16.5181 | -35.0578 | -27.1305
16.5242 | -35.0583 | -27.1386
16.5303 | -35.0589 | -27.1465
16.5364 | -35.0594 | -27.1546
16.5425 | -35.06 | -27.1626
16.5487 | -35.0605 | -27.1706
16.5548 | -35.0611 | -27.1787
16.5609 | -35.0616 | -27.1868
16.567 | -35.0622 | -27.1946
16.5731 | -35.0628 | -27.2028
16.5792 | -35.0633 | -27.2108
16.5853 | -35.0639 | -27.2189
16.5914 | -35.0644 | -27.227
16.5975 | -35.065 | -27.2351
16.6036 | -35.0655 | -27.243
16.6097 | -35.0661 | -27.2511
16.6158 | -35.0666 | -27.2591
16.6219 | -35.0672 | -27.2672
16.628 | -35.0678 | -27.2752
16.6341 | -35.0683 | -27.2832
16.6402 | -35.0689 | -27.2911
16.6463 | -35.0694 | -27.2991
16.6524 | -35.07 | -27.3071
16.6585 | -35.0705 | -27.3152
16.6646 | -35.0711 | -27.3232
16.6707 | -35.0717 | -27.3311
16.6768 | -35.0722 | -27.3392
16.6829 | -35.0728 | -27.3471
16.689 | -35.0733 | -27.3552
16.6952 | -35.0739 | -27.3631
16.7013 | -35.0744 | -27.3711
16.7074 | -35.075 | -27.3789
16.7135 | -35.0756 | -27.3869
16.7196 | -35.0761 | -27.3949
16.7257 | -35.0767 | -27.4029
16.7318 | -35.0772 | -27.4108
16.7379 | -35.0778 | -27.4188
16.744 | -35.0783 | -27.4266
16.7501 | -35.0789 | -27.4346
16.7562 | -35.0795 | -27.4426
16.7623 | -35.08 | -27.4505
16.7684 | -35.0805 | -27.4585
16.7745 | -35.0811 | -27.4666
16.7806 | -35.0817 | -27.4744
16.7867 | -35.0822 | -27.4823
16.7928 | -35.0828 | -27.4903
16.7989 | -35.0834 | -27.4982
16.805 | -35.0839 | -27.5063
16.8111 | -35.0845 | -27.5142
16.8172 | -35.085 | -27.5222
16.8233 | -35.0856 | -27.5303
16.8294 | -35.0862 | -27.5383
16.8356 | -35.0867 | -27.5464
16.8417 | -35.0873 | -27.5543
16.8478 | -35.0878 | -27.5622
16.8539 | -35.0884 | -27.5701
16.86 | -35.0889 | -27.5782
16.8661 | -35.0895 | -27.5861
16.8722 | -35.0901 | -27.5942
16.8783 | -35.0906 | -27.6022
16.8844 | -35.0912 | -27.6101
16.8905 | -35.0917 | -27.6181
16.8966 | -35.0923 | -27.626
16.9027 | -35.0929 | -27.6339
16.9088 | -35.0934 | -27.642
16.9149 | -35.094 | -27.6501
16.921 | -35.0945 | -27.6578
16.9271 | -35.0951 | -27.6659
16.9332 | -35.0957 | -27.6739
16.9393 | -35.0962 | -27.6819
16.9454 | -35.0967 | -27.6898
16.9515 | -35.0973 | -27.6978
16.9576 | -35.0979 | -27.7056
16.9637 | -35.0984 | -27.7135
16.9698 | -35.099 | -27.7214
16.9759 | -35.0995 | -27.7293
16.9821 | -35.1001 | -27.7374
16.9882 | -35.1007 | -27.7451
16.9943 | -35.1012 | -27.753
17.0004 | -35.1018 | -27.761
17.0065 | -35.1023 | -27.769
17.0126 | -35.1029 | -27.7769
17.0187 | -35.1034 | -27.785
17.0248 | -35.104 | -27.7928
17.0309 | -35.1046 | -27.8007
17.037 | -35.1051 | -27.8087
17.0431 | -35.1057 | -27.8166
17.0492 | -35.1062 | -27.8246
17.0553 | -35.1068 | -27.8324
17.0614 | -35.1073 | -27.8403
17.0675 | -35.1079 | -27.8482
17.0736 | -35.1084 | -27.8562
17.0797 | -35.109 | -27.8641
17.0858 | -35.1095 | -27.8721
17.0919 | -35.1101 | -27.88
17.098 | -35.1107 | -27.8877
17.1041 | -35.1112 | -27.8956
17.1102 | -35.1118 | -27.9035
17.1163 | -35.1124 | -27.9114
17.1225 | -35.1129 | -27.9193
17.1286 | -35.1135 | -27.9269
17.1347 | -35.1141 | -27.9349
17.1408 | -35.1146 | -27.9429
17.1469 | -35.1152 | -27.9508
17.153 | -35.1157 | -27.9587
17.1591 | -35.1163 | -27.9665
17.1652 | -35.1168 | -27.9743
17.1713 | -35.1174 | -27.9821
17.1774 | -35.1179 | -27.99
17.1835 | -35.1185 | -27.9978
17.1896 | -35.119 | -28.0057
17.1957 | -35.1196 | -28.0136
17.2018 | -35.1202 | -28.0212
17.2079 | -35.1207 | -28.0291
17.214 | -35.1213 | -28.0371
17.2201 | -35.1218 | -28.045
17.2262 | -35.1224 | -28.053
17.2323 | -35.1229 | -28.0609
17.2384 | -35.1235 | -28.0687
17.2445 | -35.1241 | -28.0765
17.2506 | -35.1246 | -28.0844
17.2567 | -35.1252 | -28.0923
17.2628 | -35.1257 | -28.1002
17.269 | -35.1263 | -28.1079
17.2751 | -35.1268 | -28.1158
17.2812 | -35.1274 | -28.1238
17.2873 | -35.1279 | -28.1317
17.2934 | -35.1285 | -28.1397
17.2995 | -35.1291 | -28.1477
17.3056 | -35.1296 | -28.1553
17.3117 | -35.1302 | -28.1633
17.3178 | -35.1307 | -28.1711
17.3239 | -35.1313 | -28.1789
17.33 | -35.1318 | -28.1868
17.3361 | -35.1324 | -28.1946
17.3422 | -35.1329 | -28.2022
17.3483 | -35.1335 | -28.2101
17.3544 | -35.134 | -28.218
17.3605 | -35.1346 | -28.2259
17.3666 | -35.1351 | -28.2337
17.3727 | -35.1357 | -28.2416
17.3788 | -35.1362 | -28.2492
17.3849 | -35.1368 | -28.257
17.391 | -35.1373 | -28.2649
17.3971 | -35.1379 | -28.2727
17.4032 | -35.1384 | -28.2805
17.4094 | -35.139 | -28.2882
17.4155 | -35.1395 | -28.296
17.4216 | -35.1401 | -28.3038
17.4277 | -35.1406 | -28.3117
17.4338 | -35.1412 | -28.3195
17.4399 | -35.1418 | -28.3274
17.446 | -35.1423 | -28.335
17.4521 | -35.1429 | -28.3428
17.4582 | -35.1434 | -28.3505
17.4643 | -35.144 | -28.3583
17.4704 | -35.1445 | -28.3661
17.4765 | -35.1451 | -28.374
17.4826 | -35.1456 | -28.3817
17.4887 | -35.1462 | -28.3896
17.4948 | -35.1467 | -28.3975
17.5009 | -35.1473 | -28.4053
17.507 | -35.1478 | -28.4132
17.5131 | -35.1484 | -28.421
17.5192 | -35.1489 | -28.4286
17.5253 | -35.1495 | -28.4364
17.5314 | -35.15 | -28.4443
17.5375 | -35.1506 | -28.4522
17.5436 | -35.1511 | -28.46
17.5497 | -35.1517 | -28.4678
17.5559 | -35.1522 | -28.4754
17.562 | -35.1528 | -28.4831
17.5681 | -35.1534 | -28.4909
17.5742 | -35.1539 | -28.4988
17.5803 | -35.1545 | -28.5065
17.5864 | -35.155 | -28.5142
17.5925 | -35.1556 | -28.5221
17.5986 | -35.1561 | -28.5298
17.6047 | -35.1567 | -28.5377
17.6108 | -35.1572 | -28.5454
17.6169 | -35.1578 | -28.5531
17.623 | -35.1584 | -28.5608
17.6291 | -35.1589 | -28.5684
17.6352 | -35.1594 | -28.5763
17.6413 | -35.16 | -28.5841
17.6474 | -35.1606 | -28.592
17.6535 | -35.1611 | -28.5997
17.6596 | -35.1617 | -28.6073
17.6657 | -35.1622 | -28.6151
17.6718 | -35.1628 | -28.623
17.6779 | -35.1633 | -28.6308
17.684 | -35.1639 | -28.6386
17.6901 | -35.1644 | -28.6464
17.6963 | -35.165 | -28.6539
17.7024 | -35.1655 | -28.6616
17.7085 | -35.1661 | -28.6693
17.7146 | -35.1666 | -28.677
17.7207 | -35.1672 | -28.6846
17.7268 | -35.1678 | -28.6922
17.7329 | -35.1683 | -28.7
17.739 | -35.1688 | -28.7078
17.7451 | -35.1694 | -28.7155
17.7512 | -35.17 | -28.7232
17.7573 | -35.1705 | -28.7311
17.7634 | -35.1711 | -28.7385
17.7695 | -35.1716 | -28.7462
17.7756 | -35.1722 | -28.754
17.7817 | -35.1727 | -28.7617
17.7878 | -35.1733 | -28.7695
17.7939 | -35.1738 | -28.7773
17.8 | -35.1744 | -28.7849
17.8061 | -35.1749 | -28.7926
17.8122 | -35.1755 | -28.8004
17.8183 | -35.176 | -28.8081
17.8244 | -35.1766 | -28.8159
17.8305 | -35.1771 | -28.8236
17.8366 | -35.1777 | -28.8311
17.8428 | -35.1782 | -28.8388
17.8489 | -35.1788 | -28.8465
17.855 | -35.1793 | -28.8541
17.8611 | -35.1799 | -28.8617
17.8672 | -35.1804 | -28.8693
17.8733 | -35.181 | -28.8771
17.8794 | -35.1815 | -28.8847
17.8855 | -35.1821 | -28.8925
17.8916 | -35.1826 | -28.9002
17.8977 | -35.1831 | -28.908
17.9038 | -35.1837 | -28.9154
17.9099 | -35.1843 | -28.9232
17.916 | -35.1848 | -28.9309
17.9221 | -35.1854 | -28.9386
17.9282 | -35.1859 | -28.9462
17.9343 | -35.1865 | -28.954
17.9404 | -35.187 | -28.9615
17.9465 | -35.1876 | -28.9692
17.9526 | -35.1881 | -28.9769
17.9587 | -35.1886 | -28.9846
17.9648 | -35.1892 | -28.9923
17.9709 | -35.1897 | -29
17.977 | -35.1903 | -29.0076
17.9832 | -35.1908 | -29.0153
17.9893 | -35.1914 | -29.0231
17.9954 | -35.1919 | -29.0309
18.0015 | -35.1925 | -29.0386
18.0076 | -35.193 | -29.0461
18.0137 | -35.1935 | -29.0538
18.0198 | -35.1941 | -29.0614
18.0259 | -35.1946 | -29.0691
18.032 | -35.1951 | -29.0769
18.0381 | -35.1957 | -29.0846
18.0442 | -35.1962 | -29.0922
18.0503 | -35.1968 | -29.0999
18.0564 | -35.1973 | -29.1076
18.0625 | -35.1979 | -29.1154
18.0686 | -35.1984 | -29.1231
18.0747 | -35.199 | -29.1309
18.0808 | -35.1995 | -29.1384
18.0869 | -35.2001 | -29.1461
18.093 | -35.2006 | -29.1537
18.0991 | -35.2012 | -29.1614
18.1052 | -35.2017 | -29.1692
18.1113 | -35.2022 | -29.1768
18.1174 | -35.2028 | -29.1844
18.1236 | -35.2034 | -29.1921
18.1297 | -35.2039 | -29.1997
18.1358 | -35.2045 | -29.2074
18.1419 | -35.205 | -29.215
18.148 | -35.2056 | -29.2227
18.1541 | -35.2061 | -29.2301
18.1602 | -35.2067 | -29.2378
18.1663 | -35.2072 | -29.2454
18.1724 | -35.2077 | -29.2531
18.1785 | -35.2083 | -29.2606
18.1846 | -35.2088 | -29.2682
18.1907 | -35.2094 | -29.2758
18.1968 | -35.2099 | -29.2835
18.2029 | -35.2105 | -29.2912
18.209 | -35.211 | -29.299
18.2151 | -35.2116 | -29.3067
18.2212 | -35.2121 | -29.3142
18.2273 | -35.2127 | -29.3218
18.2334 | -35.2132 | -29.3294
18.2395 | -35.2137 | -29.3371
18.2456 | -35.2143 | -29.3448
18.2517 | -35.2148 | -29.3524
18.2578 | -35.2154 | -29.3599
18.2639 | -35.2159 | -29.3676
18.2701 | -35.2165 | -29.3752
18.2762 | -35.217 | -29.3828
18.2823 | -35.2176 | -29.3903
18.2884 | -35.2181 | -29.3979
18.2945 | -35.2187 | -29.4052
18.3006 | -35.2192 | -29.4129
18.3067 | -35.2198 | -29.4205
18.3128 | -35.2203 | -29.4283
18.3189 | -35.2208 | -29.4359
18.325 | -35.2214 | -29.4433
18.3311 | -35.2219 | -29.451
18.3372 | -35.2225 | -29.4586
18.3433 | -35.223 | -29.4662
18.3494 | -35.2236 | -29.4739
18.3555 | -35.2241 | -29.4815
18.3616 | -35.2247 | -29.4889
18.3677 | -35.2252 | -29.4965
18.3738 | -35.2258 | -29.5041
18.3799 | -35.2263 | -29.5117
18.386 | -35.2268 | -29.5194
18.3921 | -35.2274 | -29.527
18.3982 | -35.2279 | -29.5343
18.4043 | -35.2284 | -29.5419
18.4105 | -35.229 | -29.5495
18.4166 | -35.2295 | -29.5571
18.4227 | -35.2301 | -29.5647
18.4288 | -35.2306 | -29.5723
18.4349 | -35.2311 | -29.5797
18.441 | -35.2317 | -29.5873
18.4471 | -35.2322 | -29.5947
18.4532 | -35.2327 | -29.6024
18.4593 | -35.2333 | -29.6099
18.4654 | -35.2338 | -29.6173
18.4715 | -35.2344 | -29.6249
18.4776 | -35.2349 | -29.6327
18.4837 | -35.2354 | -29.6403
18.4898 | -35.236 | -29.648
18.4959 | -35.2365 | -29.6557
18.502 | -35.237 | -29.6632
18.5081 | -35.2376 | -29.6708
18.5142 | -35.2381 | -29.6785
18.5203 | -35.2386 | -29.686
18.5264 | -35.2392 | -29.6935
18.5325 | -35.2397 | -29.7011
18.5386 | -35.2403 | -29.7086
18.5447 | -35.2408 | -29.7162
18.5508 | -35.2413 | -29.7237
18.557 | -35.2419 | -29.7313
18.5631 | -35.2424 | -29.7389
18.5692 | -35.2429 | -29.7465
18.5753 | -35.2435 | -29.7539
18.5814 | -35.244 | -29.7614
18.5875 | -35.2446 | -29.7689
18.5936 | -35.2451 | -29.7765
18.5997 | -35.2456 | -29.7842
18.6058 | -35.2462 | -29.7915
18.6119 | -35.2467 | -29.7991
18.618 | -35.2473 | -29.8067
18.6241 | -35.2478 | -29.8143
18.6302 | -35.2483 | -29.8219
18.6363 | -35.2489 | -29.8295
18.6424 | -35.2494 | -29.837
18.6485 | -35.2499 | -29.8446
18.6546 | -35.2505 | -29.8521
18.6607 | -35.251 | -29.8597
18.6668 | -35.2515 | -29.8673
18.6729 | -35.2521 | -29.875
18.679 | -35.2526 | -29.8823
18.6851 | -35.2532 | -29.8898
18.6912 | -35.2537 | -29.8972
18.6974 | -35.2542 | -29.9047
18.7035 | -35.2548 | -29.9123
18.7096 | -35.2553 | -29.9199
18.7157 | -35.2559 | -29.9273
18.7218 | -35.2564 | -29.9348
18.7279 | -35.257 | -29.9424
18.734 | -35.2575 | -29.95
18.7401 | -35.2581 | -29.9575
18.7462 | -35.2586 | -29.9648
18.7523 | -35.2592 | -29.9724
18.7584 | -35.2597 | -29.98
18.7645 | -35.2603 | -29.9875
18.7706 | -35.2608 | -29.995
18.7767 | -35.2613 | -30.0025
18.7828 | -35.2619 | -30.0098
18.7889 | -35.2624 | -30.0174
18.795 | -35.2629 | -30.0251
18.8011 | -35.2635 | -30.0326
18.8072 | -35.264 | -30.0402
18.8133 | -35.2646 | -30.0477
18.8194 | -35.2651 | -30.0551
18.8255 | -35.2656 | -30.0626
18.8316 | -35.2662 | -30.0701
18.8377 | -35.2667 | -30.0776
18.8439 | -35.2673 | -30.085
18.85 | -35.2678 | -30.0926
18.8561 | -35.2683 | -30.0999
18.8622 | -35.2689 | -30.1075
18.8683 | -35.2694 | -30.1148
18.8744 | -35.2699 | -30.1223
18.8805 | -35.2705 | -30.1299
18.8866 | -35.271 | -30.1375
18.8927 | -35.2715 | -30.1448
18.8988 | -35.2721 | -30.1522
18.9049 | -35.2726 | -30.1597
18.911 | -35.2732 | -30.1672
18.9171 | -35.2737 | -30.1746
18.9232 | -35.2743 | -30.1819
18.9293 | -35.2748 | -30.1894
18.9354 | -35.2754 | -30.1969
18.9415 | -35.2759 | -30.2044
18.9476 | -35.2764 | -30.2118
18.9537 | -35.277 | -30.2193
18.9598 | -35.2775 | -30.2265
18.9659 | -35.278 | -30.234
18.972 | -35.2786 | -30.2414
18.9781 | -35.2791 | -30.249
18.9843 | -35.2796 | -30.2566
18.9904 | -35.2802 | -30.264
18.9965 | -35.2807 | -30.2713
19.0026 | -35.2813 | -30.2788
19.0087 | -35.2818 | -30.2863
19.0148 | -35.2824 | -30.2938
19.0209 | -35.2829 | -30.3013
19.027 | -35.2834 | -30.3089
19.0331 | -35.284 | -30.3161
19.0392 | -35.2845 | -30.3235
19.0453 | -35.2851 | -30.331
19.0514 | -35.2856 | -30.3383
19.0575 | -35.2861 | -30.3459
19.0636 | -35.2867 | -30.353
19.0697 | -35.2872 | -30.3605
19.0758 | -35.2877 | -30.3679
19.0819 | -35.2883 | -30.3753
19.088 | -35.2888 | -30.3827
19.0941 | -35.2893 | -30.3903
19.1002 | -35.2899 | -30.3976
19.1063 | -35.2904 | -30.405
19.1124 | -35.2909 | -30.4125
19.1185 | -35.2915 | -30.4199
19.1246 | -35.292 | -30.4273
19.1308 | -35.2926 | -30.4348
19.1369 | -35.2931 | -30.4421
19.143 | -35.2936 | -30.4496
19.1491 | -35.2942 | -30.4571
19.1552 | -35.2947 | -30.4647
19.1613 | -35.2952 | -30.4721
19.1674 | -35.2958 | -30.4795
19.1735 | -35.2963 | -30.4868
19.1796 | -35.2968 | -30.4942
19.1857 | -35.2974 | -30.5017
19.1918 | -35.2979 | -30.5091
19.1979 | -35.2984 | -30.5166
19.204 | -35.299 | -30.5238
19.2101 | -35.2995 | -30.5312
19.2162 | -35.3001 | -30.5386
19.2223 | -35.3006 | -30.5459
19.2284 | -35.3011 | -30.5534
19.2345 | -35.3017 | -30.5608
19.2406 | -35.3022 | -30.568
19.2467 | -35.3027 | -30.5755
19.2528 | -35.3033 | -30.583
19.2589 | -35.3038 | -30.5903
19.265 | -35.3043 | -30.5977
19.2712 | -35.3049 | -30.6051
19.2773 | -35.3054 | -30.6124
19.2834 | -35.306 | -30.6197
19.2895 | -35.3065 | -30.6272
19.2956 | -35.307 | -30.6346
19.3017 | -35.3075 | -30.642
19.3078 | -35.3081 | -30.6495
19.3139 | -35.3086 | -30.6567
19.32 | -35.3092 | -30.664
19.3261 | -35.3097 | -30.6713
19.3322 | -35.3102 | -30.6787
19.3383 | -35.3107 | -30.6861
19.3444 | -35.3113 | -30.6932
19.3505 | -35.3118 | -30.7006
19.3566 | -35.3123 | -30.708
19.3627 | -35.3128 | -30.7154
19.3688 | -35.3134 | -30.7228
19.3749 | -35.3139 | -30.7302
19.381 | -35.3145 | -30.7374
19.3871 | -35.315 | -30.7446
19.3932 | -35.3155 | -30.752
19.3993 | -35.3161 | -30.7594
19.4054 | -35.3166 | -30.7669
19.4115 | -35.3171 | -30.7743
19.4177 | -35.3177 | -30.7815
19.4238 | -35.3182 | -30.7889
19.4299 | -35.3187 | -30.7962
19.436 | -35.3193 | -30.8037
19.4421 | -35.3198 | -30.8111
19.4482 | -35.3203 | -30.8185
19.4543 | -35.3208 | -30.8255
19.4604 | -35.3214 | -30.8329
19.4665 | -35.3219 | -30.8403
19.4726 | -35.3224 | -30.8475
19.4787 | -35.3229 | -30.8549
19.4848 | -35.3235 | -30.8622
19.4909 | -35.324 | -30.8694
19.497 | -35.3245 | -30.8768
19.5031 | -35.3251 | -30.884
19.5092 | -35.3256 | -30.8914
19.5153 | -35.3261 | -30.8987
19.5214 | -35.3266 | -30.9058
19.5275 | -35.3272 | -30.9131
19.5336 | -35.3277 | -30.9204
19.5397 | -35.3282 | -30.9277
19.5458 | -35.3288 | -30.935
19.5519 | -35.3293 | -30.9423
19.5581 | -35.3298 | -30.9494
19.5642 | -35.3303 | -30.9567
19.5703 | -35.3309 | -30.9641
19.5764 | -35.3314 | -30.9713
19.5825 | -35.3319 | -30.9786
19.5886 | -35.3325 | -30.986
19.5947 | -35.333 | -30.9931
19.6008 | -35.3335 | -31.0004
19.6069 | -35.3341 | -31.0077
19.613 | -35.3346 | -31.0151
19.6191 | -35.3351 | -31.0223
19.6252 | -35.3356 | -31.0296
19.6313 | -35.3362 | -31.0366
19.6374 | -35.3367 | -31.044
19.6435 | -35.3372 | -31.0512
19.6496 | -35.3377 | -31.0585
19.6557 | -35.3383 | -31.0658
19.6618 | -35.3388 | -31.0728
19.6679 | -35.3393 | -31.0801
19.674 | -35.3399 | -31.0874
19.6801 | -35.3404 | -31.0947
19.6862 | -35.3409 | -31.1021
19.6923 | -35.3415 | -31.1092
19.6984 | -35.342 | -31.1164
19.7046 | -35.3426 | -31.1237
19.7107 | -35.3431 | -31.1309
19.7168 | -35.3436 | -31.1382
19.7229 | -35.3441 | -31.1455
19.729 | -35.3446 | -31.1528
19.7351 | -35.3452 | -31.1598
19.7412 | -35.3457 | -31.1673
19.7473 | -35.3463 | -31.1745
19.7534 | -35.3468 | -31.1819
19.7595 | -35.3473 | -31.1893
19.7656 | -35.3478 | -31.1966
19.7717 | -35.3484 | -31.2035
19.7778 | -35.3489 | -31.2107
19.7839 | -35.3494 | -31.2179
19.79 | -35.3499 | -31.2252
19.7961 | -35.3505 | -31.2325
19.8022 | -35.351 | -31.2395
19.8083 | -35.3515 | -31.2467
19.8144 | -35.3521 | -31.2539
19.8205 | -35.3526 | -31.2611
19.8266 | -35.3531 | -31.2682
19.8327 | -35.3537 | -31.2754
19.8388 | -35.3542 | -31.2824
19.845 | -35.3547 | -31.2896
19.8511 | -35.3552 | -31.297
19.8572 | -35.3557 | -31.3042
19.8633 | -35.3563 | -31.3114
19.8694 | -35.3568 | -31.3188
19.8755 | -35.3573 | -31.3258
19.8816 | -35.3578 | -31.3332
19.8877 | -35.3584 | -31.3406
19.8938 | -35.3589 | -31.3478
19.8999 | -35.3594 | -31.355
19.906 | -35.3599 | -31.3622
19.9121 | -35.3605 | -31.3693
19.9182 | -35.361 | -31.3766
19.9243 | -35.3615 | -31.3838
19.9304 | -35.362 | -31.391
19.9365 | -35.3625 | -31.3982
19.9426 | -35.3631 | -31.4053
19.9487 | -35.3636 | -31.4125
19.9548 | -35.3641 | -31.4198
19.9609 | -35.3647 | -31.427
19.967 | -35.3652 | -31.4341
19.9731 | -35.3657 | -31.4413
19.9792 | -35.3663 | -31.4484
19.9853 | -35.3668 | -31.4556
19.9915 | -35.3673 | -31.4627
19.9976 | -35.3678 | -31.4699
20.0037 | -35.3684 | -31.4771
20.0098 | -35.3689 | -31.4843
20.0159 | -35.3694 | -31.4913
20.022 | -35.37 | -31.4985
20.0281 | -35.3705 | -31.5058
20.0342 | -35.371 | -31.513
20.0403 | -35.3716 | -31.5202
20.0464 | -35.3721 | -31.5274
20.0525 | -35.3726 | -31.5343
20.0586 | -35.3732 | -31.5415
20.0647 | -35.3737 | -31.5487
20.0708 | -35.3742 | -31.5561
20.0769 | -35.3747 | -31.5633
20.083 | -35.3753 | -31.5702
20.0891 | -35.3758 | -31.5775
20.0952 | -35.3763 | -31.5846
20.1013 | -35.3768 | -31.5919
20.1074 | -35.3774 | -31.599
20.1135 | -35.3779 | -31.6062
20.1196 | -35.3784 | -31.6131
20.1257 | -35.3789 | -31.6203
20.1319 | -35.3795 | -31.6275
20.138 | -35.38 | -31.6347
20.1441 | -35.3805 | -31.6419
20.1502 | -35.3811 | -31.6492
20.1563 | -35.3816 | -31.6561
20.1624 | -35.3821 | -31.6633
20.1685 | -35.3826 | -31.6706
20.1746 | -35.3832 | -31.6777
20.1807 | -35.3837 | -31.6848
20.1868 | -35.3842 | -31.692
20.1929 | -35.3847 | -31.6989
20.199 | -35.3853 | -31.7061
20.2051 | -35.3858 | -31.7131
20.2112 | -35.3863 | -31.7203
20.2173 | -35.3868 | -31.7274
20.2234 | -35.3874 | -31.7345
20.2295 | -35.3879 | -31.7414
20.2356 | -35.3885 | -31.7486
20.2417 | -35.389 | -31.7557
20.2478 | -35.3895 | -31.7629
20.2539 | -35.39 | -31.77
20.26 | -35.3906 | -31.7769
20.2661 | -35.3911 | -31.7841
20.2723 | -35.3916 | -31.7912
20.2784 | -35.3922 | -31.7983
20.2845 | -35.3927 | -31.8054
20.2906 | -35.3932 | -31.8124
20.2967 | -35.3938 | -31.8194
20.3028 | -35.3943 | -31.8265
20.3089 | -35.3948 | -31.8336
20.315 | -35.3953 | -31.8406
20.3211 | -35.3959 | -31.8477
20.3272 | -35.3964 | -31.8549
20.3333 | -35.3969 | -31.8617
20.3394 | -35.3975 | -31.8688
20.3455 | -35.398 | -31.8759
20.3516 | -35.3985 | -31.8828
20.3577 | -35.399 | -31.8899
20.3638 | -35.3995 | -31.8969
20.3699 | -35.4001 | -31.9038
20.376 | -35.4006 | -31.9108
20.3821 | -35.4011 | -31.918
20.3882 | -35.4017 | -31.925
20.3943 | -35.4022 | -31.9321
20.4004 | -35.4027 | -31.9389
20.4065 | -35.4032 | -31.9459
20.4126 | -35.4038 | -31.953
20.4188 | -35.4043 | -31.9601
20.4249 | -35.4048 | -31.9671
20.431 | -35.4053 | -31.9741
20.4371 | -35.4059 | -31.981
20.4432 | -35.4064 | -31.9881
20.4493 | -35.4069 | -31.9952
20.4554 | -35.4074 | -32.0022
20.4615 | -35.408 | -32.0093
20.4676 | -35.4085 | -32.0165
20.4737 | -35.409 | -32.0233
20.4798 | -35.4095 | -32.0304
20.4859 | -35.4101 | -32.0376
20.492 | -35.4106 | -32.0446
20.4981 | -35.4111 | -32.0516
20.5042 | -35.4116 | -32.0587
20.5103 | -35.4122 | -32.0655
20.5164 | -35.4127 | -32.0725
20.5225 | -35.4132 | -32.0795
20.5286 | -35.4137 | -32.0866
20.5347 | -35.4143 | -32.0935
20.5408 | -35.4148 | -32.1003
20.5469 | -35.4153 | -32.1074
20.553 | -35.4158 | -32.1145
20.5592 | -35.4164 | -32.1215
20.5653 | -35.4169 | -32.1285
20.5714 | -35.4174 | -32.1356
20.5775 | -35.418 | -32.1424
20.5836 | -35.4185 | -32.1495
20.5897 | -35.419 | -32.1566
20.5958 | -35.4195 | -32.1637
20.6019 | -35.4201 | -32.1708
20.608 | -35.4206 | -32.1779
20.6141 | -35.4211 | -32.1847
20.6202 | -35.4217 | -32.1917
20.6263 | -35.4222 | -32.1987
20.6324 | -35.4227 | -32.2056
20.6385 | -35.4233 | -32.2126
20.6446 | -35.4238 | -32.2197
20.6507 | -35.4243 | -32.2265
20.6568 | -35.4249 | -32.2335
20.6629 | -35.4254 | -32.2405
20.669 | -35.4259 | -32.2476
20.6751 | -35.4264 | -32.2546
20.6812 | -35.427 | -32.2613
20.6873 | -35.4275 | -32.2684
20.6934 | -35.428 | -32.2754
20.6995 | -35.4285 | -32.2823
20.7057 | -35.4291 | -32.2894
20.7118 | -35.4296 | -32.2965
20.7179 | -35.4301 | -32.3032
20.724 | -35.4307 | -32.3103
20.7301 | -35.4312 | -32.3172
20.7362 | -35.4317 | -32.3243
20.7423 | -35.4322 | -32.3313
20.7484 | -35.4328 | -32.3383
20.7545 | -35.4333 | -32.345
20.7606 | -35.4339 | -32.352
20.7667 | -35.4344 | -32.3589
20.7728 | -35.435 | -32.3659
20.7789 | -35.4355 | -32.3728
20.785 | -35.436 | -32.3798
20.7911 | -35.4365 | -32.3866
20.7972 | -35.4371 | -32.3936
20.8033 | -35.4376 | -32.4007
20.8094 | -35.4381 | -32.4076
20.8155 | -35.4386 | -32.4145
20.8216 | -35.4392 | -32.4215
20.8277 | -35.4397 | -32.4281
20.8338 | -35.4402 | -32.4352
20.8399 | -35.4408 | -32.4419
20.8461 | -35.4413 | -32.4489
20.8522 | -35.4418 | -32.4559
20.8583 | -35.4424 | -32.4626
20.8644 | -35.4429 | -32.4695
20.8705 | -35.4434 | -32.4766
20.8766 | -35.4439 | -32.4837
20.8827 | -35.4445 | -32.4906
20.8888 | -35.445 | -32.4975
20.8949 | -35.4455 | -32.5042
20.901 | -35.446 | -32.5112
20.9071 | -35.4466 | -32.5181
20.9132 | -35.4471 | -32.5249
20.9193 | -35.4476 | -32.5318
20.9254 | -35.4482 | -32.5388
20.9315 | -35.4487 | -32.5455
20.9376 | -35.4492 | -32.5525
20.9437 | -35.4497 | -32.5594
20.9498 | -35.4503 | -32.5663
20.9559 | -35.4508 | -32.5732
20.962 | -35.4513 | -32.58
20.9681 | -35.4519 | -32.5869
20.9742 | -35.4524 | -32.5938
20.9803 | -35.4529 | -32.6008
20.9864 | -35.4535 | -32.6077
20.9926 | -35.454 | -32.6148
20.9987 | -35.4546 | -32.6214
21.0048 | -35.4551 | -32.6282
21.0109 | -35.4556 | -32.6352
21.017 | -35.4562 | -32.6421
21.0231 | -35.4567 | -32.649
21.0292 | -35.4572 | -32.6559
21.0353 | -35.4578 | -32.6626
21.0414 | -35.4583 | -32.6694
21.0475 | -35.4588 | -32.6763
21.0536 | -35.4593 | -32.6832
21.0597 | -35.4599 | -32.6901
21.0658 | -35.4604 | -32.697
21.0719 | -35.4609 | -32.704
21.078 | -35.4614 | -32.7108
21.0841 | -35.4619 | -32.7179
21.0902 | -35.4624 | -32.7243
21.0963 | -35.463 | -32.7308
21.1024 | -35.4635 | -32.7378
21.1085 | -35.4641 | -32.7445
21.1146 | -35.4646 | -32.7514
21.1207 | -35.4651 | -32.7582
21.1268 | -35.4657 | -32.7652
21.133 | -35.4662 | -32.7721
21.1391 | -35.4668 | -32.7786
21.1452 | -35.4673 | -32.7854
21.1513 | -35.4678 | -32.7922
21.1574 | -35.4683 | -32.7991
21.1635 | -35.4689 | -32.806
21.1696 | -35.4694 | -32.8127
21.1757 | -35.4699 | -32.8193
21.1818 | -35.4704 | -32.8262
21.1879 | -35.471 | -32.833
21.194 | -35.4715 | -32.8398
21.2001 | -35.472 | -32.8466
21.2062 | -35.4726 | -32.8535
21.2123 | -35.4731 | -32.8601
21.2184 | -35.4736 | -32.8668
21.2245 | -35.4742 | -32.8737
21.2306 | -35.4747 | -32.8804
21.2367 | -35.4752 | -32.8872
21.2428 | -35.4757 | -32.894
21.2489 | -35.4763 | -32.9005
21.255 | -35.4768 | -32.9075
21.2611 | -35.4773 | -32.9144
21.2672 | -35.4779 | -32.9213
21.2733 | -35.4784 | -32.9282
21.2795 | -35.4789 | -32.9346
21.2856 | -35.4795 | -32.9413
21.2917 | -35.48 | -32.9481
21.2978 | -35.4805 | -32.955
21.3039 | -35.4811 | -32.9618
21.31 | -35.4816 | -32.9687
21.3161 | -35.4821 | -32.9752
21.3222 | -35.4827 | -32.9821
21.3283 | -35.4832 | -32.9889
21.3344 | -35.4837 | -32.9957
21.3405 | -35.4842 | -33.0025
21.3466 | -35.4848 | -33.0093
21.3527 | -35.4853 | -33.0159
21.3588 | -35.4858 | -33.0227
21.3649 | -35.4864 | -33.0294
21.371 | -35.4869 | -33.0362
21.3771 | -35.4874 | -33.043
21.3832 | -35.488 | -33.0498
21.3893 | -35.4885 | -33.0564
21.3954 | -35.489 | -33.0632
21.4015 | -35.4896 | -33.0699
21.4076 | -35.4901 | -33.0768
21.4137 | -35.4906 | -33.0836
21.4199 | -35.4911 | -33.0905
21.426 | -35.4917 | -33.097
21.4321 | -35.4922 | -33.1037
21.4382 | -35.4927 | -33.1106
21.4443 | -35.4932 | -33.1172
21.4504 | -35.4937 | -33.1241
21.4565 | -35.4943 | -33.1304
21.4626 | -35.4948 | -33.1372
21.4687 | -35.4953 | -33.144
21.4748 | -35.4959 | -33.1507
21.4809 | -35.4964 | -33.1574
21.487 | -35.4969 | -33.1641
21.4931 | -35.4974 | -33.1707
21.4992 | -35.498 | -33.1775
21.5053 | -35.4985 | -33.1843
21.5114 | -35.499 | -33.1911
21.5175 | -35.4996 | -33.1978
21.5236 | -35.5001 | -33.2046
21.5297 | -35.5006 | -33.211
21.5358 | -35.5011 | -33.2178
21.5419 | -35.5017 | -33.2246
21.548 | -35.5022 | -33.2313
21.5541 | -35.5027 | -33.2381
21.5602 | -35.5033 | -33.2448
21.5664 | -35.5038 | -33.2514
21.5725 | -35.5043 | -33.2582
21.5786 | -35.5048 | -33.2649
21.5847 | -35.5054 | -33.2718
21.5908 | -35.5059 | -33.2785
21.5969 | -35.5064 | -33.2848
21.603 | -35.507 | -33.2916
21.6091 | -35.5075 | -33.2984
21.6152 | -35.508 | -33.3052
21.6213 | -35.5085 | -33.3119
21.6274 | -35.5091 | -33.3187
21.6335 | -35.5096 | -33.3251
21.6396 | -35.5101 | -33.3319
21.6457 | -35.5106 | -33.3387
21.6518 | -35.5112 | -33.3453
21.6579 | -35.5117 | -33.3521
21.664 | -35.5122 | -33.359
21.6701 | -35.5128 | -33.3655
21.6762 | -35.5133 | -33.3722
21.6823 | -35.5138 | -33.3789
21.6884 | -35.5143 | -33.3856
21.6945 | -35.5149 | -33.3923
21.7006 | -35.5154 | -33.3991
21.7068 | -35.5159 | -33.4054
21.7129 | -35.5165 | -33.4121
21.719 | -35.517 | -33.4188
21.7251 | -35.5175 | -33.4255
21.7312 | -35.518 | -33.4322
21.7373 | -35.5186 | -33.4385
21.7434 | -35.5191 | -33.4453
21.7495 | -35.5197 | -33.452
21.7556 | -35.5202 | -33.4587
21.7617 | -35.5207 | -33.4654
21.7678 | -35.5212 | -33.472
21.7739 | -35.5218 | -33.4783
21.78 | -35.5223 | -33.485
21.7861 | -35.5229 | -33.4918
21.7922 | -35.5234 | -33.4986
21.7983 | -35.5239 | -33.5052
21.8044 | -35.5244 | -33.5119
21.8105 | -35.525 | -33.5183
21.8166 | -35.5255 | -33.5249
21.8227 | -35.5261 | -33.5316
21.8288 | -35.5266 | -33.5384
21.8349 | -35.5271 | -33.5451
21.841 | -35.5276 | -33.5517
21.8471 | -35.5282 | -33.558
21.8533 | -35.5287 | -33.5647
21.8594 | -35.5293 | -33.5714
21.8655 | -35.5298 | -33.578
21.8716 | -35.5303 | -33.5847
21.8777 | -35.5309 | -33.5913
21.8838 | -35.5314 | -33.5979
21.8899 | -35.5319 | -33.6045
21.896 | -35.5325 | -33.6112
21.9021 | -35.533 | -33.6179
21.9082 | -35.5336 | -33.6246
21.9143 | -35.5341 | -33.631
21.9204 | -35.5346 | -33.6376
21.9265 | -35.5352 | -33.6442
21.9326 | -35.5357 | -33.6508
21.9387 | -35.5362 | -33.6575
21.9448 | -35.5367 | -33.6642
21.9509 | -35.5373 | -33.6707
21.957 | -35.5378 | -33.6774
21.9631 | -35.5383 | -33.684
21.9692 | -35.5389 | -33.6907
21.9753 | -35.5394 | -33.6973
21.9814 | -35.5399 | -33.704
21.9875 | -35.5405 | -33.7103
21.9937 | -35.541 | -33.717
21.9998 | -35.5415 | -33.7236
22.0059 | -35.5421 | -33.7303
22.012 | -35.5426 | -33.7369
22.0181 | -35.5431 | -33.7433
22.0242 | -35.5437 | -33.7498
22.0303 | -35.5442 | -33.7565
22.0364 | -35.5447 | -33.763
22.0425 | -35.5453 | -33.7696
22.0486 | -35.5458 | -33.7762
22.0547 | -35.5464 | -33.7826
22.0608 | -35.5469 | -33.7892
22.0669 | -35.5474 | -33.7957
22.073 | -35.548 | -33.8022
22.0791 | -35.5485 | -33.8088
22.0852 | -35.549 | -33.8154
22.0913 | -35.5495 | -33.8218
22.0974 | -35.5501 | -33.8283
22.1035 | -35.5506 | -33.8349
22.1096 | -35.5511 | -33.8415
22.1157 | -35.5516 | -33.8482
22.1218 | -35.5522 | -33.8548
22.1279 | -35.5527 | -33.8611
22.134 | -35.5532 | -33.8677
22.1402 | -35.5538 | -33.8742
22.1463 | -35.5543 | -33.8808
22.1524 | -35.5548 | -33.8875
22.1585 | -35.5554 | -33.8941
22.1646 | -35.5559 | -33.9005
22.1707 | -35.5564 | -33.9071
22.1768 | -35.557 | -33.9137
22.1829 | -35.5575 | -33.9203
22.189 | -35.558 | -33.9269
22.1951 | -35.5585 | -33.9332
22.2012 | -35.5591 | -33.9398
22.2073 | -35.5596 | -33.9464
22.2134 | -35.5601 | -33.9528
22.2195 | -35.5607 | -33.9593
22.2256 | -35.5612 | -33.966
22.2317 | -35.5618 | -33.9723
22.2378 | -35.5623 | -33.979
22.2439 | -35.5628 | -33.9855
22.25 | -35.5633 | -33.992
22.2561 | -35.5639 | -33.9985
22.2622 | -35.5644 | -34.005
22.2683 | -35.5649 | -34.0113
22.2744 | -35.5655 | -34.0178
22.2806 | -35.566 | -34.0244
22.2867 | -35.5666 | -34.0309
22.2928 | -35.5671 | -34.0374
22.2989 | -35.5676 | -34.0439
22.305 | -35.5682 | -34.0502
22.3111 | -35.5687 | -34.0568
22.3172 | -35.5693 | -34.0634
22.3233 | -35.5698 | -34.0699
22.3294 | -35.5704 | -34.0765
22.3355 | -35.5709 | -34.0828
22.3416 | -35.5714 | -34.0894
22.3477 | -35.572 | -34.0959
22.3538 | -35.5725 | -34.1024
22.3599 | -35.573 | -34.1089
22.366 | -35.5735 | -34.1153
22.3721 | -35.5741 | -34.1217
22.3782 | -35.5746 | -34.1283
22.3843 | -35.5751 | -34.1348
22.3904 | -35.5757 | -34.1413
22.3965 | -35.5762 | -34.1478
22.4026 | -35.5767 | -34.1544
22.4087 | -35.5773 | -34.1606
22.4148 | -35.5778 | -34.1671
22.4209 | -35.5783 | -34.1736
22.4271 | -35.5789 | -34.1801
22.4332 | -35.5794 | -34.1866
22.4393 | -35.5799 | -34.1931
22.4454 | -35.5805 | -34.1994
22.4515 | -35.581 | -34.2059
22.4576 | -35.5815 | -34.2124
22.4637 | -35.5821 | -34.2189
22.4698 | -35.5826 | -34.2254
22.4759 | -35.5832 | -34.2317
22.482 | -35.5837 | -34.2382
22.4881 | -35.5842 | -34.2447
22.4942 | -35.5847 | -34.2513
22.5003 | -35.5853 | -34.2579
22.5064 | -35.5858 | -34.2643
22.5125 | -35.5863 | -34.2705
22.5186 | -35.5869 | -34.277
22.5247 | -35.5874 | -34.2835
22.5308 | -35.5879 | -34.29
22.5369 | -35.5884 | -34.2966
22.543 | -35.5889 | -34.3031
22.5491 | -35.5895 | -34.3094
22.5552 | -35.59 | -34.3158
22.5613 | -35.5906 | -34.3224
22.5675 | -35.5911 | -34.3289
22.5736 | -35.5916 | -34.3354
22.5797 | -35.5922 | -34.3418
22.5858 | -35.5927 | -34.348
22.5919 | -35.5933 | -34.3545
22.598 | -35.5938 | -34.361
22.6041 | -35.5944 | -34.3675
22.6102 | -35.5949 | -34.374
22.6163 | -35.5954 | -34.3802
22.6224 | -35.596 | -34.3867
22.6285 | -35.5965 | -34.3932
22.6346 | -35.597 | -34.3997
22.6407 | -35.5976 | -34.4061
22.6468 | -35.5981 | -34.4125
22.6529 | -35.5987 | -34.4187
22.659 | -35.5992 | -34.4251
22.6651 | -35.5997 | -34.4316
22.6712 | -35.6003 | -34.438
22.6773 | -35.6008 | -34.4445
22.6834 | -35.6013 | -34.451
22.6895 | -35.6019 | -34.4572
22.6956 | -35.6024 | -34.4636
22.7017 | -35.6029 | -34.4701
22.7079 | -35.6035 | -34.4764
22.714 | -35.604 | -34.483
22.7201 | -35.6045 | -34.4894
22.7262 | -35.6051 | -34.4957
22.7323 | -35.6056 | -34.5021
22.7384 | -35.6061 | -34.5086
22.7445 | -35.6067 | -34.5151
22.7506 | -35.6072 | -34.5215
22.7567 | -35.6077 | -34.5279
22.7628 | -35.6083 | -34.5341
22.7689 | -35.6088 | -34.5406
22.775 | -35.6094 | -34.5471
22.7811 | -35.6099 | -34.5535
22.7872 | -35.6104 | -34.5599
22.7933 | -35.611 | -34.5661
22.7994 | -35.6115 | -34.5725
22.8055 | -35.612 | -34.579
22.8116 | -35.6126 | -34.5853
22.8177 | -35.6131 | -34.5919
22.8238 | -35.6136 | -34.5982
22.8299 | -35.6142 | -34.6044
22.836 | -35.6147 | -34.6108
22.8421 | -35.6153 | -34.6174
22.8482 | -35.6158 | -34.6238
22.8544 | -35.6163 | -34.6302
22.8605 | -35.6169 | -34.6366
22.8666 | -35.6174 | -34.6429
22.8727 | -35.6179 | -34.6492
22.8788 | -35.6185 | -34.6555
22.8849 | -35.619 | -34.6618
22.891 | -35.6195 | -34.6682
22.8971 | -35.62 | -34.6747
22.9032 | -35.6206 | -34.6809
22.9093 | -35.6211 | -34.6873
22.9154 | -35.6216 | -34.6936
22.9215 | -35.6222 | -34.7001
22.9276 | -35.6227 | -34.7065
22.9337 | -35.6232 | -34.7125
22.9398 | -35.6238 | -34.7189
22.9459 | -35.6243 | -34.7254
22.952 | -35.6249 | -34.7318
22.9581 | -35.6254 | -34.7382
22.9642 | -35.6259 | -34.7447
22.9703 | -35.6265 | -34.7508
22.9764 | -35.627 | -34.7572
22.9825 | -35.6275 | -34.7637
22.9886 | -35.628 | -34.7701
22.9948 | -35.6286 | -34.7764
23.0009 | -35.6291 | -34.7828
23.007 | -35.6296 | -34.7891
23.0131 | -35.6302 | -34.7954
23.0192 | -35.6307 | -34.8017
23.0253 | -35.6312 | -34.8082
23.0314 | -35.6318 | -34.8146
23.0375 | -35.6323 | -34.821
23.0436 | -35.6328 | -34.8271
23.0497 | -35.6334 | -34.8335
23.0558 | -35.6339 | -34.8399
23.0619 | -35.6344 | -34.8463
23.068 | -35.635 | -34.8527
23.0741 | -35.6355 | -34.8587
23.0802 | -35.636 | -34.865
23.0863 | -35.6365 | -34.8714
23.0924 | -35.6371 | -34.8777
23.0985 | -35.6376 | -34.884
23.1046 | -35.6381 | -34.8904
23.1107 | -35.6387 | -34.8967
23.1168 | -35.6392 | -34.903
23.1229 | -35.6397 | -34.9094
23.129 | -35.6403 | -34.9157
23.1351 | -35.6408 | -34.9219
23.1413 | -35.6413 | -34.9282
23.1474 | -35.6419 | -34.9344
23.1535 | -35.6424 | -34.9407
23.1596 | -35.6429 | -34.947
23.1657 | -35.6435 | -34.9534
23.1718 | -35.644 | -34.9597
23.1779 | -35.6445 | -34.9661
23.184 | -35.6451 | -34.9722
23.1901 | -35.6456 | -34.9786
23.1962 | -35.6462 | -34.9849
23.2023 | -35.6467 | -34.9913
23.2084 | -35.6472 | -34.9977
23.2145 | -35.6478 | -35.0037
23.2206 | -35.6483 | -35.0101
23.2267 | -35.6488 | -35.0165
23.2328 | -35.6494 | -35.0227
23.2389 | -35.6499 | -35.0291
23.245 | -35.6504 | -35.0354
23.2511 | -35.651 | -35.0414
23.2572 | -35.6515 | -35.0478
23.2633 | -35.652 | -35.0541
23.2694 | -35.6526 | -35.0605
23.2755 | -35.6531 | -35.0669
23.2817 | -35.6536 | -35.0732
23.2878 | -35.6542 | -35.0794
23.2939 | -35.6547 | -35.0857
23.3 | -35.6552 | -35.0921
23.3061 | -35.6558 | -35.0985
23.3122 | -35.6563 | -35.1048
23.3183 | -35.6568 | -35.111
23.3244 | -35.6574 | -35.1172
23.3305 | -35.6579 | -35.1235
23.3366 | -35.6584 | -35.1298
23.3427 | -35.6589 | -35.136
23.3488 | -35.6595 | -35.1423
23.3549 | -35.66 | -35.1486
23.361 | -35.6606 | -35.1548
23.3671 | -35.6611 | -35.1611
23.3732 | -35.6616 | -35.1674
23.3793 | -35.6621 | -35.1736
23.3854 | -35.6627 | -35.1801
23.3915 | -35.6632 | -35.1862
23.3976 | -35.6638 | -35.1925
23.4037 | -35.6643 | -35.1987
23.4098 | -35.6649 | -35.2049
23.4159 | -35.6654 | -35.2113
23.422 | -35.6659 | -35.2176
23.4282 | -35.6665 | -35.2237
23.4343 | -35.667 | -35.23
23.4404 | -35.6675 | -35.2364
23.4465 | -35.6681 | -35.2427
23.4526 | -35.6686 | -35.249
23.4587 | -35.6691 | -35.2553
23.4648 | -35.6696 | -35.2613
23.4709 | -35.6702 | -35.2675
23.477 | -35.6707 | -35.2739
23.4831 | -35.6712 | -35.2801
23.4892 | -35.6717 | -35.2863
23.4953 | -35.6723 | -35.2926
23.5014 | -35.6728 | -35.2987
23.5075 | -35.6733 | -35.3051
23.5136 | -35.6739 | -35.3113
23.5197 | -35.6744 | -35.3175
23.5258 | -35.6749 | -35.3238
23.5319 | -35.6754 | -35.3299
23.538 | -35.676 | -35.3361
23.5441 | -35.6765 | -35.3424
23.5502 | -35.677 | -35.3486
23.5563 | -35.6775 | -35.355
23.5624 | -35.6781 | -35.3613
23.5686 | -35.6786 | -35.3673
23.5747 | -35.6791 | -35.3735
23.5808 | -35.6797 | -35.3798
23.5869 | -35.6802 | -35.386
23.593 | -35.6807 | -35.3922
23.5991 | -35.6813 | -35.3986
23.6052 | -35.6818 | -35.4048
23.6113 | -35.6823 | -35.4111
23.6174 | -35.6828 | -35.4174
23.6235 | -35.6833 | -35.4237
23.6296 | -35.6839 | -35.43
23.6357 | -35.6844 | -35.4362
23.6418 | -35.6849 | -35.4423
23.6479 | -35.6854 | -35.4485
23.654 | -35.686 | -35.4547
23.6601 | -35.6865 | -35.461
23.6662 | -35.687 | -35.4672
23.6723 | -35.6876 | -35.4731
23.6784 | -35.6881 | -35.4794
23.6845 | -35.6886 | -35.4856
23.6906 | -35.6891 | -35.4917
23.6967 | -35.6897 | -35.4979
23.7028 | -35.6902 | -35.5043
23.7089 | -35.6907 | -35.5103
23.7151 | -35.6913 | -35.5165
23.7212 | -35.6918 | -35.5229
23.7273 | -35.6923 | -35.5292
23.7334 | -35.6928 | -35.5354
23.7395 | -35.6933 | -35.5416
23.7456 | -35.6938 | -35.5476
23.7517 | -35.6944 | -35.5538
23.7578 | -35.6949 | -35.5599
23.7639 | -35.6954 | -35.5662
23.77 | -35.6959 | -35.5725
23.7761 | -35.6965 | -35.5788
23.7822 | -35.697 | -35.5849
23.7883 | -35.6975 | -35.5911
23.7944 | -35.698 | -35.5973
23.8005 | -35.6985 | -35.6036
23.8066 | -35.6991 | -35.6099
23.8127 | -35.6996 | -35.616
23.8188 | -35.7001 | -35.6224
23.8249 | -35.7006 | -35.6286
23.831 | -35.7011 | -35.6349
23.8371 | -35.7017 | -35.641
23.8432 | -35.7022 | -35.6472
23.8493 | -35.7027 | -35.6532
23.8555 | -35.7032 | -35.6594
23.8616 | -35.7037 | -35.6656
23.8677 | -35.7042 | -35.6719
23.8738 | -35.7047 | -35.678
23.8799 | -35.7053 | -35.6842
23.886 | -35.7058 | -35.6901
23.8921 | -35.7063 | -35.6963
23.8982 | -35.7068 | -35.7025
23.9043 | -35.7073 | -35.7088
23.9104 | -35.7079 | -35.715
23.9165 | -35.7084 | -35.7212
23.9226 | -35.7089 | -35.7273
23.9287 | -35.7094 | -35.7336
23.9348 | -35.71 | -35.7398
23.9409 | -35.7105 | -35.7461
23.947 | -35.711 | -35.7523
23.9531 | -35.7115 | -35.7582
23.9592 | -35.712 | -35.7644
23.9653 | -35.7126 | -35.7706
23.9714 | -35.7131 | -35.7767
23.9775 | -35.7136 | -35.7829
23.9836 | -35.7141 | -35.789
23.9897 | -35.7147 | -35.795
23.9958 | -35.7152 | -35.8012
24.002 | -35.7157 | -35.8074
24.0081 | -35.7162 | -35.8135
24.0142 | -35.7167 | -35.8197
24.0203 | -35.7173 | -35.8259
24.0264 | -35.7178 | -35.832
24.0325 | -35.7183 | -35.8382
24.0386 | -35.7188 | -35.8444
24.0447 | -35.7193 | -35.8505
24.0508 | -35.7198 | -35.8567
24.0569 | -35.7204 | -35.8629
24.063 | -35.7209 | -35.8687
24.0691 | -35.7214 | -35.8748
24.0752 | -35.722 | -35.8811
24.0813 | -35.7225 | -35.8872
24.0874 | -35.723 | -35.8934
24.0935 | -35.7235 | -35.8996
24.0996 | -35.724 | -35.9055
24.1057 | -35.7245 | -35.9117
24.1118 | -35.725 | -35.918
24.1179 | -35.7255 | -35.9242
24.124 | -35.7261 | -35.9303
24.1301 | -35.7266 | -35.9362
24.1362 | -35.7271 | -35.9423
24.1424 | -35.7276 | -35.9485
24.1485 | -35.7281 | -35.9548
24.1546 | -35.7286 | -35.9608
24.1607 | -35.7291 | -35.967
24.1668 | -35.7297 | -35.9728
24.1729 | -35.7302 | -35.979
24.179 | -35.7307 | -35.9851
24.1851 | -35.7312 | -35.9912
24.1912 | -35.7317 | -35.9973
24.1973 | -35.7322 | -36.0035
24.2034 | -35.7327 | -36.0094
24.2095 | -35.7333 | -36.0157
24.2156 | -35.7338 | -36.0218
24.2217 | -35.7343 | -36.028
24.2278 | -35.7348 | -36.0341
24.2339 | -35.7353 | -36.0401
24.24 | -35.7358 | -36.046
24.2461 | -35.7363 | -36.052
24.2522 | -35.7368 | -36.0582
24.2583 | -35.7373 | -36.0644
24.2644 | -35.7379 | -36.0706
24.2705 | -35.7384 | -36.0767
24.2766 | -35.7389 | -36.0827
24.2827 | -35.7394 | -36.0888
24.2889 | -35.7399 | -36.0949
24.295 | -35.7404 | -36.101
24.3011 | -35.7409 | -36.1071
24.3072 | -35.7415 | -36.113
24.3133 | -35.742 | -36.1191
24.3194 | -35.7425 | -36.1253
24.3255 | -35.743 | -36.1315
24.3316 | -35.7435 | -36.1376
24.3377 | -35.744 | -36.1438
24.3438 | -35.7445 | -36.1497
24.3499 | -35.745 | -36.1557
24.356 | -35.7455 | -36.1619
24.3621 | -35.746 | -36.168
24.3682 | -35.7465 | -36.1741
24.3743 | -35.747 | -36.1803
24.3804 | -35.7475 | -36.1862
24.3865 | -35.748 | -36.1922
24.3926 | -35.7485 | -36.1984
24.3987 | -35.749 | -36.2046
24.4048 | -35.7495 | -36.2107
24.4109 | -35.75 | -36.2168
24.417 | -35.7505 | -36.2229
24.4231 | -35.751 | -36.229
24.4293 | -35.7516 | -36.235
24.4354 | -35.7521 | -36.2411
24.4415 | -35.7526 | -36.2471
24.4476 | -35.7531 | -36.2531
24.4537 | -35.7536 | -36.2592
24.4598 | -35.7541 | -36.2652
24.4659 | -35.7546 | -36.2714
24.472 | -35.7551 | -36.2776
24.4781 | -35.7556 | -36.2837
24.4842 | -35.7561 | -36.2896
24.4903 | -35.7566 | -36.2957
24.4964 | -35.7571 | -36.3018
24.5025 | -35.7576 | -36.3078
24.5086 | -35.7582 | -36.3139
24.5147 | -35.7587 | -36.32
24.5208 | -35.7592 | -36.3258
24.5269 | -35.7597 | -36.332
24.533 | -35.7602 | -36.3381
24.5391 | -35.7607 | -36.3442
24.5452 | -35.7612 | -36.3503
24.5513 | -35.7617 | -36.3561
24.5574 | -35.7622 | -36.3622
24.5635 | -35.7627 | -36.3683
24.5696 | -35.7632 | -36.3744
24.5758 | -35.7638 | -36.3804
24.5819 | -35.7643 | -36.3864
24.588 | -35.7647 | -36.3924
24.5941 | -35.7652 | -36.3986
24.6002 | -35.7657 | -36.4046
24.6063 | -35.7663 | -36.4108
24.6124 | -35.7668 | -36.4168
24.6185 | -35.7673 | -36.4229
24.6246 | -35.7678 | -36.4289
24.6307 | -35.7683 | -36.435
24.6368 | -35.7688 | -36.4411
24.6429 | -35.7693 | -36.4471
24.649 | -35.7698 | -36.4533
24.6551 | -35.7703 | -36.4593
24.6612 | -35.7708 | -36.4652
24.6673 | -35.7713 | -36.4713
24.6734 | -35.7718 | -36.4774
24.6795 | -35.7723 | -36.4836
24.6856 | -35.7728 | -36.4897
24.6917 | -35.7733 | -36.4957
24.6978 | -35.7738 | -36.5015
24.7039 | -35.7743 | -36.5076
24.71 | -35.7748 | -36.5137
24.7162 | -35.7753 | -36.5199
24.7223 | -35.7759 | -36.5259
24.7284 | -35.7764 | -36.5318
24.7345 | -35.7769 | -36.5379
24.7406 | -35.7774 | -36.5439
24.7467 | -35.7779 | -36.55
24.7528 | -35.7784 | -36.556
24.7589 | -35.7789 | -36.5621
24.765 | -35.7793 | -36.5679
24.7711 | -35.7799 | -36.5739
24.7772 | -35.7804 | -36.5799
24.7833 | -35.7808 | -36.586
24.7894 | -35.7814 | -36.5921
24.7955 | -35.7819 | -36.5982
24.8016 | -35.7823 | -36.6041
24.8077 | -35.7828 | -36.6101
24.8138 | -35.7833 | -36.6163
24.8199 | -35.7838 | -36.6223
24.826 | -35.7843 | -36.6283
24.8321 | -35.7848 | -36.6342
24.8382 | -35.7853 | -36.64
24.8443 | -35.7858 | -36.6461
24.8504 | -35.7863 | -36.6521
24.8565 | -35.7868 | -36.6582
24.8627 | -35.7873 | -36.6642
24.8688 | -35.7878 | -36.6702
24.8749 | -35.7883 | -36.6762
24.881 | -35.7888 | -36.6823
24.8871 | -35.7893 | -36.6884
24.8932 | -35.7898 | -36.6944
24.8993 | -35.7903 | -36.7005
24.9054 | -35.7908 | -36.7064
24.9115 | -35.7913 | -36.7124
24.9176 | -35.7918 | -36.7185
24.9237 | -35.7923 | -36.7245
24.9298 | -35.7928 | -36.7305
24.9359 | -35.7933 | -36.7367
24.942 | -35.7937 | -36.7424
24.9481 | -35.7942 | -36.7484
24.9542 | -35.7947 | -36.7545
24.9603 | -35.7952 | -36.7605
24.9664 | -35.7957 | -36.7665
24.9725 | -35.7962 | -36.7725
24.9786 | -35.7967 | -36.7782
24.9847 | -35.7972 | -36.7841
24.9908 | -35.7977 | -36.7902
24.9969 | -35.7981 | -36.7963
25.0031 | -35.7987 | -36.8023
25.0092 | -35.7991 | -36.8081
25.0153 | -35.7996 | -36.814
25.0214 | -35.8001 | -36.8201
25.0275 | -35.8006 | -36.826
25.0336 | -35.8011 | -36.8319
25.0397 | -35.8016 | -36.838
25.0458 | -35.8021 | -36.8438
25.0519 | -35.8025 | -36.8498
25.058 | -35.803 | -36.8558
25.0641 | -35.8035 | -36.8618
25.0702 | -35.804 | -36.8677
25.0763 | -35.8045 | -36.8736
25.0824 | -35.805 | -36.8794
25.0885 | -35.8055 | -36.8854
25.0946 | -35.806 | -36.8914
25.1007 | -35.8064 | -36.8975
25.1068 | -35.8069 | -36.9034
25.1129 | -35.8074 | -36.9094
25.119 | -35.8079 | -36.9152
25.1251 | -35.8084 | -36.9211
25.1312 | -35.8089 | -36.9272
25.1373 | -35.8094 | -36.9332
25.1435 | -35.8099 | -36.9392
25.1496 | -35.8104 | -36.945
25.1557 | -35.8109 | -36.9509
25.1618 | -35.8113 | -36.9569
25.1679 | -35.8118 | -36.963
25.174 | -35.8123 | -36.969
25.1801 | -35.8128 | -36.9751
25.1862 | -35.8133 | -36.981
25.1923 | -35.8138 | -36.987
25.1984 | -35.8143 | -36.993
25.2045 | -35.8147 | -36.999
25.2106 | -35.8152 | -37.005
25.2167 | -35.8157 | -37.0111
25.2228 | -35.8162 | -37.0169
25.2289 | -35.8166 | -37.0229
25.235 | -35.8171 | -37.0289
25.2411 | -35.8176 | -37.0349
25.2472 | -35.8181 | -37.0409
25.2533 | -35.8186 | -37.047
25.2594 | -35.819 | -37.0528
25.2655 | -35.8195 | -37.0587
25.2716 | -35.82 | -37.0646
25.2777 | -35.8205 | -37.0705
25.2838 | -35.821 | -37.0766
25.29 | -35.8215 | -37.0825
25.2961 | -35.8219 | -37.0882
25.3022 | -35.8224 | -37.0941
25.3083 | -35.8229 | -37.1001
25.3144 | -35.8234 | -37.106
25.3205 | -35.8239 | -37.112
25.3266 | -35.8244 | -37.1178
25.3327 | -35.8248 | -37.1238
25.3388 | -35.8253 | -37.1298
25.3449 | -35.8258 | -37.1358
25.351 | -35.8263 | -37.1417
25.3571 | -35.8268 | -37.1477
25.3632 | -35.8272 | -37.1535
25.3693 | -35.8277 | -37.1595
25.3754 | -35.8282 | -37.1654
25.3815 | -35.8287 | -37.1714
25.3876 | -35.8291 | -37.1774
25.3937 | -35.8296 | -37.1833
25.3998 | -35.8301 | -37.1891
25.4059 | -35.8305 | -37.1951
25.412 | -35.831 | -37.201
25.4181 | -35.8315 | -37.2068
25.4242 | -35.832 | -37.2128
25.4304 | -35.8325 | -37.2188
25.4365 | -35.8329 | -37.2245
25.4426 | -35.8334 | -37.2304
25.4487 | -35.8339 | -37.2365
25.4548 | -35.8344 | -37.2424
25.4609 | -35.8348 | -37.2483
25.467 | -35.8353 | -37.254
25.4731 | -35.8358 | -37.26
25.4792 | -35.8363 | -37.2659
25.4853 | -35.8368 | -37.2719
25.4914 | -35.8372 | -37.2778
25.4975 | -35.8377 | -37.2837
25.5036 | -35.8382 | -37.2895
25.5097 | -35.8387 | -37.2954
25.5158 | -35.8392 | -37.3013
25.5219 | -35.8396 | -37.3073
25.528 | -35.8401 | -37.3132
25.5341 | -35.8406 | -37.3191
25.5402 | -35.8411 | -37.3249
25.5463 | -35.8416 | -37.3308
25.5524 | -35.842 | -37.3368
25.5585 | -35.8425 | -37.3426
25.5646 | -35.843 | -37.3486
25.5707 | -35.8435 | -37.3545
25.5769 | -35.8439 | -37.3601
25.583 | -35.8444 | -37.3661
25.5891 | -35.8449 | -37.3721
25.5952 | -35.8454 | -37.378
25.6013 | -35.8459 | -37.3839
25.6074 | -35.8463 | -37.3896
25.6135 | -35.8468 | -37.3955
25.6196 | -35.8473 | -37.4016
25.6257 | -35.8477 | -37.4074
25.6318 | -35.8482 | -37.4133
25.6379 | -35.8487 | -37.4192
25.644 | -35.8491 | -37.4249
25.6501 | -35.8496 | -37.4307
25.6562 | -35.8501 | -37.4366
25.6623 | -35.8505 | -37.4423
25.6684 | -35.851 | -37.4481
25.6745 | -35.8515 | -37.454
25.6806 | -35.852 | -37.4597
25.6867 | -35.8524 | -37.4656
25.6928 | -35.8529 | -37.4715
25.6989 | -35.8534 | -37.4775
25.705 | -35.8538 | -37.4834
25.7111 | -35.8543 | -37.4894
25.7173 | -35.8548 | -37.4951
25.7234 | -35.8553 | -37.501
25.7295 | -35.8558 | -37.5068
25.7356 | -35.8562 | -37.5128
25.7417 | -35.8567 | -37.5186
25.7478 | -35.8571 | -37.5244
25.7539 | -35.8576 | -37.5302
25.76 | -35.8581 | -37.5361
25.7661 | -35.8585 | -37.5419
25.7722 | -35.859 | -37.5478
25.7783 | -35.8595 | -37.5537
25.7844 | -35.8599 | -37.5593
25.7905 | -35.8604 | -37.5652
25.7966 | -35.8609 | -37.5711
25.8027 | -35.8613 | -37.577
25.8088 | -35.8618 | -37.583
25.8149 | -35.8622 | -37.5889
25.821 | -35.8627 | -37.5947
25.8271 | -35.8632 | -37.6005
25.8332 | -35.8636 | -37.6064
25.8393 | -35.8641 | -37.6122
25.8454 | -35.8646 | -37.6181
25.8515 | -35.865 | -37.6241
25.8576 | -35.8655 | -37.6299
25.8638 | -35.866 | -37.6358
25.8699 | -35.8664 | -37.6416
25.876 | -35.8669 | -37.6474
25.8821 | -35.8674 | -37.6532
25.8882 | -35.8678 | -37.6589
25.8943 | -35.8683 | -37.6649
25.9004 | -35.8687 | -37.6706
25.9065 | -35.8692 | -37.6766
25.9126 | -35.8697 | -37.6825
25.9187 | -35.8701 | -37.6883
25.9248 | -35.8706 | -37.6939
25.9309 | -35.8711 | -37.6997
25.937 | -35.8715 | -37.7055
25.9431 | -35.872 | -37.7114
25.9492 | -35.8724 | -37.7172
25.9553 | -35.8729 | -37.7231
25.9614 | -35.8734 | -37.7288
25.9675 | -35.8739 | -37.7346
25.9736 | -35.8743 | -37.7406
25.9797 | -35.8748 | -37.7464
25.9858 | -35.8753 | -37.7523
25.9919 | -35.8757 | -37.7582
25.998 | -35.8761 | -37.764
26.0042 | -35.8766 | -37.7698
26.0103 | -35.8771 | -37.7757
26.0164 | -35.8775 | -37.7815
26.0225 | -35.878 | -37.7874
26.0286 | -35.8784 | -37.7931
26.0347 | -35.8789 | -37.7987
26.0408 | -35.8793 | -37.8046
26.0469 | -35.8798 | -37.8103
26.053 | -35.8802 | -37.8162
26.0591 | -35.8807 | -37.8221
26.0652 | -35.8811 | -37.8277
26.0713 | -35.8816 | -37.8334
26.0774 | -35.882 | -37.8391
26.0835 | -35.8825 | -37.845
26.0896 | -35.8829 | -37.8509
26.0957 | -35.8834 | -37.8567
26.1018 | -35.8838 | -37.8623
26.1079 | -35.8843 | -37.8682
26.114 | -35.8847 | -37.874
26.1201 | -35.8852 | -37.8797
26.1262 | -35.8857 | -37.8855
26.1323 | -35.8861 | -37.8913
26.1384 | -35.8866 | -37.8969
26.1445 | -35.887 | -37.9027
26.1507 | -35.8875 | -37.9086
26.1568 | -35.8879 | -37.9144
26.1629 | -35.8884 | -37.9202
26.169 | -35.8888 | -37.926
26.1751 | -35.8893 | -37.9316
26.1812 | -35.8897 | -37.9374
26.1873 | -35.8902 | -37.9432
26.1934 | -35.8907 | -37.949
26.1995 | -35.8911 | -37.9548
26.2056 | -35.8916 | -37.9605
26.2117 | -35.892 | -37.9664
26.2178 | -35.8925 | -37.9721
26.2239 | -35.8929 | -37.9779
26.23 | -35.8934 | -37.9837
26.2361 | -35.8939 | -37.9895
26.2422 | -35.8943 | -37.9952
26.2483 | -35.8947 | -38.0009
26.2544 | -35.8952 | -38.0066
26.2605 | -35.8956 | -38.0124
26.2666 | -35.8961 | -38.0182
26.2727 | -35.8965 | -38.0239
26.2788 | -35.897 | -38.0295
26.2849 | -35.8974 | -38.0353
26.2911 | -35.8979 | -38.0411
26.2972 | -35.8984 | -38.0469
26.3033 | -35.8988 | -38.0526
26.3094 | -35.8993 | -38.0584
26.3155 | -35.8997 | -38.0639
26.3216 | -35.9002 | -38.0698
26.3277 | -35.9006 | -38.0755
26.3338 | -35.9011 | -38.0812
26.3399 | -35.9016 | -38.0868
26.346 | -35.902 | -38.0923
26.3521 | -35.9025 | -38.0982
26.3582 | -35.9029 | -38.1041
26.3643 | -35.9034 | -38.1098
26.3704 | -35.9038 | -38.1155
26.3765 | -35.9043 | -38.1212
26.3826 | -35.9047 | -38.1266
26.3887 | -35.9052 | -38.1323
26.3948 | -35.9056 | -38.1381
26.4009 | -35.9061 | -38.144
26.407 | -35.9066 | -38.1497
26.4131 | -35.907 | -38.1555
26.4192 | -35.9075 | -38.1609
26.4253 | -35.9079 | -38.1667
26.4314 | -35.9084 | -38.1724
26.4376 | -35.9088 | -38.1781
26.4437 | -35.9093 | -38.1838
26.4498 | -35.9097 | -38.1895
26.4559 | -35.9102 | -38.195
26.462 | -35.9106 | -38.2008
26.4681 | -35.911 | -38.2065
26.4742 | -35.9115 | -38.2123
26.4803 | -35.912 | -38.218
26.4864 | -35.9124 | -38.2235
26.4925 | -35.9128 | -38.2291
26.4986 | -35.9133 | -38.2349
26.5047 | -35.9138 | -38.2406
26.5108 | -35.9142 | -38.2463
26.5169 | -35.9147 | -38.252
26.523 | -35.9151 | -38.2576
26.5291 | -35.9156 | -38.2632
26.5352 | -35.916 | -38.2688
26.5413 | -35.9164 | -38.2746
26.5474 | -35.9169 | -38.2804
26.5535 | -35.9173 | -38.2861
26.5596 | -35.9177 | -38.2917
26.5657 | -35.9182 | -38.2975
26.5718 | -35.9186 | -38.3031
26.578 | -35.9191 | -38.3089
26.5841 | -35.9195 | -38.3144
26.5902 | -35.92 | -38.3202
26.5963 | -35.9204 | -38.3257
26.6024 | -35.9208 | -38.3314
26.6085 | -35.9213 | -38.3372
26.6146 | -35.9217 | -38.3429
26.6207 | -35.9222 | -38.3487
26.6268 | -35.9226 | -38.3545
26.6329 | -35.923 | -38.36
26.639 | -35.9235 | -38.3656
26.6451 | -35.9239 | -38.3714
26.6512 | -35.9244 | -38.3771
26.6573 | -35.9248 | -38.3829
26.6634 | -35.9253 | -38.3883
26.6695 | -35.9257 | -38.3941
26.6756 | -35.9262 | -38.3998
26.6817 | -35.9266 | -38.4055
26.6878 | -35.927 | -38.4112
26.6939 | -35.9275 | -38.4169
26.7 | -35.9279 | -38.4224
26.7061 | -35.9283 | -38.4279
26.7122 | -35.9288 | -38.4336
26.7183 | -35.9292 | -38.4393
26.7245 | -35.9296 | -38.445
26.7306 | -35.9301 | -38.4507
26.7367 | -35.9305 | -38.4562
26.7428 | -35.931 | -38.4618
26.7489 | -35.9314 | -38.4675
26.755 | -35.9318 | -38.4731
26.7611 | -35.9323 | -38.4788
26.7672 | -35.9327 | -38.4846
26.7733 | -35.9331 | -38.4902
26.7794 | -35.9336 | -38.4959
26.7855 | -35.9341 | -38.5016
26.7916 | -35.9345 | -38.507
26.7977 | -35.9349 | -38.5127
26.8038 | -35.9354 | -38.5182
26.8099 | -35.9358 | -38.5239
26.816 | -35.9362 | -38.5296
26.8221 | -35.9367 | -38.5353
26.8282 | -35.9371 | -38.541
26.8343 | -35.9376 | -38.5466
26.8404 | -35.938 | -38.5521
26.8465 | -35.9384 | -38.5578
26.8526 | -35.9389 | -38.5634
26.8587 | -35.9393 | -38.569
26.8649 | -35.9398 | -38.5746
26.871 | -35.9402 | -38.5803
26.8771 | -35.9406 | -38.5857
26.8832 | -35.941 | -38.5913
26.8893 | -35.9415 | -38.5969
26.8954 | -35.9419 | -38.6025
26.9015 | -35.9423 | -38.6082
26.9076 | -35.9428 | -38.6138
26.9137 | -35.9432 | -38.6192
26.9198 | -35.9436 | -38.625
26.9259 | -35.9441 | -38.6305
26.932 | -35.9445 | -38.636
26.9381 | -35.9449 | -38.6417
26.9442 | -35.9454 | -38.6472
26.9503 | -35.9458 | -38.6527
26.9564 | -35.9462 | -38.6583
26.9625 | -35.9467 | -38.664
26.9686 | -35.9471 | -38.6696
26.9747 | -35.9475 | -38.6752
26.9808 | -35.948 | -38.6806
26.9869 | -35.9484 | -38.6861
26.993 | -35.9489 | -38.6918
26.9991 | -35.9493 | -38.6973
27.0052 | -35.9497 | -38.703
27.0114 | -35.9502 | -38.7088
27.0175 | -35.9506 | -38.7142
27.0236 | -35.951 | -38.7199
27.0297 | -35.9514 | -38.7254
27.0358 | -35.9519 | -38.7311
27.0419 | -35.9523 | -38.7368
27.048 | -35.9528 | -38.7425
27.0541 | -35.9532 | -38.7478
27.0602 | -35.9536 | -38.7534
27.0663 | -35.9541 | -38.7591
27.0724 | -35.9545 | -38.7646
27.0785 | -35.9549 | -38.7703
27.0846 | -35.9554 | -38.7758
27.0907 | -35.9558 | -38.7814
27.0968 | -35.9562 | -38.7871
27.1029 | -35.9567 | -38.7927
27.109 | -35.9571 | -38.7982
27.1151 | -35.9575 | -38.8039
27.1212 | -35.9579 | -38.8093
27.1273 | -35.9584 | -38.8149
27.1334 | -35.9588 | -38.8205
27.1395 | -35.9592 | -38.8262
27.1456 | -35.9597 | -38.8318
27.1518 | -35.9601 | -38.8374
27.1579 | -35.9605 | -38.8427
27.164 | -35.961 | -38.8483
27.1701 | -35.9614 | -38.8539
27.1762 | -35.9618 | -38.8594
27.1823 | -35.9622 | -38.865
27.1884 | -35.9626 | -38.8706
27.1945 | -35.9631 | -38.8759
27.2006 | -35.9635 | -38.8814
27.2067 | -35.9639 | -38.8869
27.2128 | -35.9644 | -38.8925
27.2189 | -35.9648 | -38.8982
27.225 | -35.9652 | -38.9038
27.2311 | -35.9657 | -38.9092
27.2372 | -35.9661 | -38.9147
27.2433 | -35.9666 | -38.9201
27.2494 | -35.967 | -38.9257
27.2555 | -35.9675 | -38.9312
27.2616 | -35.9679 | -38.9366
27.2677 | -35.9683 | -38.9422
27.2738 | -35.9688 | -38.9478
27.2799 | -35.9692 | -38.9535
27.286 | -35.9696 | -38.9589
27.2921 | -35.97 | -38.9645
27.2983 | -35.9704 | -38.9698
27.3044 | -35.9709 | -38.9754
27.3105 | -35.9713 | -38.981
27.3166 | -35.9717 | -38.9865
27.3227 | -35.9722 | -38.992
27.3288 | -35.9726 | -38.9975
27.3349 | -35.973 | -39.0028
27.341 | -35.9734 | -39.0083
27.3471 | -35.9739 | -39.0138
27.3532 | -35.9743 | -39.0193
27.3593 | -35.9747 | -39.0249
27.3654 | -35.9751 | -39.0305
27.3715 | -35.9756 | -39.0357
27.3776 | -35.976 | -39.0413
27.3837 | -35.9764 | -39.0467
27.3898 | -35.9768 | -39.0522
27.3959 | -35.9772 | -39.0577
27.402 | -35.9777 | -39.0631
27.4081 | -35.9781 | -39.0686
27.4142 | -35.9785 | -39.0742
27.4203 | -35.979 | -39.0797
27.4264 | -35.9794 | -39.0852
27.4325 | -35.9798 | -39.0907
27.4387 | -35.9802 | -39.096
27.4448 | -35.9807 | -39.1015
27.4509 | -35.9811 | -39.1071
27.457 | -35.9815 | -39.1128
27.4631 | -35.9819 | -39.1183
27.4692 | -35.9824 | -39.1238
27.4753 | -35.9828 | -39.129
27.4814 | -35.9832 | -39.1346
27.4875 | -35.9836 | -39.1401
27.4936 | -35.984 | -39.1456
27.4997 | -35.9845 | -39.1511
27.5058 | -35.9849 | -39.1566
27.5119 | -35.9853 | -39.162
27.518 | -35.9857 | -39.1675
27.5241 | -35.9862 | -39.1729
27.5302 | -35.9866 | -39.1784
27.5363 | -35.987 | -39.184
27.5424 | -35.9874 | -39.1893
27.5485 | -35.9878 | -39.1949
27.5546 | -35.9883 | -39.2003
27.5607 | -35.9887 | -39.2058
27.5668 | -35.9891 | -39.2113
27.5729 | -35.9895 | -39.2168
27.5791 | -35.99 | -39.2221
27.5852 | -35.9904 | -39.2276
27.5913 | -35.9908 | -39.233
27.5974 | -35.9912 | -39.2385
27.6035 | -35.9917 | -39.244
27.6096 | -35.9921 | -39.2495
27.6157 | -35.9925 | -39.2547
27.6218 | -35.9929 | -39.2603
27.6279 | -35.9933 | -39.2657
27.634 | -35.9938 | -39.2712
27.6401 | -35.9942 | -39.2767
27.6462 | -35.9946 | -39.2822
27.6523 | -35.995 | -39.2873
27.6584 | -35.9954 | -39.2927
27.6645 | -35.9959 | -39.2983
27.6706 | -35.9963 | -39.3038
27.6767 | -35.9967 | -39.3092
27.6828 | -35.9971 | -39.3145
27.6889 | -35.9975 | -39.3201
27.695 | -35.9979 | -39.3254
27.7011 | -35.9983 | -39.3309
27.7072 | -35.9988 | -39.3363
27.7133 | -35.9992 | -39.3418
27.7194 | -35.9996 | -39.3472
27.7256 | -36.0001 | -39.3528
27.7317 | -36.0005 | -39.3582
27.7378 | -36.0009 | -39.3637
27.7439 | -36.0013 | -39.3692
27.75 | -36.0018 | -39.3747
27.7561 | -36.0022 | -39.38
27.7622 | -36.0026 | -39.3854
27.7683 | -36.003 | -39.3908
27.7744 | -36.0035 | -39.3962
27.7805 | -36.0039 | -39.4017
27.7866 | -36.0043 | -39.4072
27.7927 | -36.0047 | -39.4123
27.7988 | -36.0052 | -39.4177
27.8049 | -36.0056 | -39.4232
27.811 | -36.006 | -39.4287
27.8171 | -36.0064 | -39.4342
27.8232 | -36.0069 | -39.4393
27.8293 | -36.0073 | -39.4448
27.8354 | -36.0077 | -39.4502
27.8415 | -36.0081 | -39.4557
27.8476 | -36.0085 | -39.4612
27.8537 | -36.0089 | -39.4667
27.8598 | -36.0094 | -39.4719
27.866 | -36.0098 | -39.4773
27.8721 | -36.0102 | -39.4829
27.8782 | -36.0106 | -39.4884
27.8843 | -36.011 | -39.4937
27.8904 | -36.0115 | -39.4992
27.8965 | -36.0119 | -39.5044
27.9026 | -36.0123 | -39.5099
27.9087 | -36.0127 | -39.5153
27.9148 | -36.0131 | -39.5207
27.9209 | -36.0136 | -39.5261
27.927 | -36.014 | -39.5314
27.9331 | -36.0144 | -39.5367
27.9392 | -36.0148 | -39.5421
27.9453 | -36.0152 | -39.5476
27.9514 | -36.0156 | -39.553
27.9575 | -36.0161 | -39.5585
27.9636 | -36.0165 | -39.564
27.9697 | -36.0169 | -39.5692
27.9758 | -36.0173 | -39.5748
27.9819 | -36.0177 | -39.5802
27.988 | -36.0181 | -39.5855
27.9941 | -36.0185 | -39.5909
28.0002 | -36.019 | -39.5961
28.0063 | -36.0194 | -39.6016
28.0125 | -36.0198 | -39.6071
28.0186 | -36.0202 | -39.6125
28.0247 | -36.0206 | -39.6179
28.0308 | -36.021 | -39.6232
28.0369 | -36.0215 | -39.6284
28.043 | -36.0219 | -39.6338
28.0491 | -36.0223 | -39.639
28.0552 | -36.0227 | -39.6445
28.0613 | -36.0231 | -39.6498
28.0674 | -36.0235 | -39.6551
28.0735 | -36.024 | -39.6603

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.7121 | -44.0218
21.4908 | -35.6508 | -36.3438
16.4948 | -34.9725 | -28.6798
12.6602 | -34.578 | -22.163
9.7171 | -34.3407 | -17.082
7.4582 | -34.1806 | -12.9604
5.7244 | -34.0721 | -9.6195
4.3936 | -34.0082 | -6.9291
3.3722 | -33.9853 | -4.8172
2.5883 | -33.983 | -3.2075
1.9866 | -33.9827 | -1.9247
1.5248 | -33.9845 | -0.89859
1.1703 | -33.9968 | -0.22376
1 | -33.9911 | 0.11305
0.89822 | -34.0047 | 0.38784
0.68941 | -33.977 | 0.98471
0.52919 | -33.9892 | 1.4515
0.40612 | -33.9868 | 2.0689
0.31171 | -33.9742 | 2.5264
0.23928 | -33.9774 | 3.2518
0.18358 | -33.9999 | 3.6594
0.14091 | -34.0172 | 4.535
0.10819 | -34.0111 | 5.7091
0.083017 | -34.0597 | 7.1641
0.063705 | -34.1004 | 9.0366
0.048923 | -34.1664 | 11.5112
0.037479 | -34.2663 | 14.6935
0.028801 | -34.4389 | 18.7144
0.022078 | -34.7562 | 24.1508
0.017023 | -35.2392 | 30.5707
0.013018 | -36.042 | 38.2268
0.010014 | -37.1533 | 46.1392

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.00357 | -2.7448
0.00714 | -31.1256
0.01071 | -36.7167
0.01428 | -41.4148
0.01785 | -44.6866
0.02142 | -46.0074
0.02499 | -47.2923
0.02856 | -48.6432
0.03213 | -49.9849
0.0357 | -51.0804
0.03927 | -53.9243
0.04284 | -54.8348
0.04641 | -54.2058
0.04998 | -55.0747
0.05355 | -57.4626
0.05712 | -56.7226
0.06069 | -56.988
0.06426 | -57.5455
0.06783 | -57.7528
0.0714 | -58.7817
0.07497 | -59.5895
0.07854 | -60.0196
0.08211 | -61.5384
0.08568 | -61.0151
0.08925 | -62.033
0.09282 | -61.4
0.09639 | -60.9752
0.09996 | -62.1728
0.10353 | -63.2744
0.1071 | -63.2964
0.11067 | -63.668
0.11424 | -63.9641
0.11781 | -64.4364
0.12138 | -64.2933
0.12495 | -64.8437
0.12852 | -65.2428
0.13209 | -65.1972
0.13566 | -65.4428
0.13923 | -65.1529
0.1428 | -66.2652
0.14637 | -66.9126
0.14994 | -67.8064
0.15351 | -67.4118
0.15708 | -67.0507
0.16065 | -67.0828
0.16422 | -67.4344
0.16779 | -68.3506
0.17136 | -67.6359
0.17493 | -67.9797
0.1785 | -68.2937
0.18207 | -68.9737
0.18564 | -68.6932
0.18921 | -68.4475
0.19278 | -69.5267
0.19635 | -69.8445
0.19992 | -69.835
0.20349 | -70.3893
0.20706 | -70.6951
0.21063 | -71.7362
0.2142 | -71.2426
0.21777 | -70.9945
0.22134 | -71.8222
0.22491 | -72.0216
0.22848 | -70.1787
0.23205 | -70.068
0.23562 | -71.6219
0.23919 | -71.442
0.24276 | -71.9569
0.24633 | -72.1915
0.2499 | -72.1371
0.25347 | -72.5607
0.25704 | -73.0014
0.26061 | -72.4017
0.26418 | -71.9186
0.26775 | -72.1646
0.27132 | -72.1033
0.27489 | -72.7597
0.27846 | -73.6901
0.28203 | -73.2433
0.2856 | -73.4949
0.28917 | -73.3879
0.29274 | -72.2927
0.29631 | -72.9647
0.29988 | -74.5857
0.30345 | -73.7986
0.30702 | -73.7203
0.31059 | -74.2197
0.31416 | -74.2258
0.31773 | -74.5994
0.3213 | -73.4027
0.32487 | -72.9341
0.32844 | -74.5242
0.33201 | -74.9121
0.33558 | -74.9469
0.33915 | -74.3401
0.34272 | -74.6581
0.34629 | -75.2466
0.34986 | -74.8308
0.35343 | -74.2539
0.357 | -75.3044
0.36057 | -75.36
0.36414 | -74.7747
0.36771 | -74.8291
0.37128 | -75.2658
0.37485 | -75.371
0.37842 | -74.9826
0.38199 | -76.8097
0.38556 | -76.546
0.38913 | -75.2792
0.3927 | -75.669
0.39627 | -76.7488
0.39984 | -76.1621
0.40341 | -76.317
0.40698 | -77.333
0.41055 | -75.9413
0.41412 | -75.83
0.41769 | -76.9366
0.42126 | -76.8524
0.42483 | -75.9425
0.4284 | -75.9304
0.43197 | -76.4881
0.43554 | -76.9179
0.43911 | -77.564
0.44268 | -78.2221
0.44625 | -78.7579
0.44982 | -78.4209
0.45339 | -77.5299
0.45696 | -77.8117
0.46053 | -77.5536
0.4641 | -77.3587
0.46767 | -79.1406
0.47124 | -78.7203
0.47481 | -76.8682
0.47838 | -77.3129
0.48195 | -78.2436
0.48552 | -78.1476
0.48909 | -78.5451
0.49266 | -77.5918
0.49623 | -77.7345
0.4998 | -78.261
0.50337 | -79.1447
0.50694 | -79.6931
0.51051 | -78.7063
0.51408 | -78.823
0.51765 | -79.1878
0.52122 | -79.4283
0.52479 | -79.3995
0.52836 | -79.2237
0.53193 | -79.2338
0.5355 | -79.2748
0.53907 | -79.2491
0.54264 | -79.3533
0.54621 | -79.0011
0.54978 | -78.7363
0.55335 | -79.1163
0.55692 | -79.3111
0.56049 | -79.6124
0.56406 | -80.4052
0.56763 | -80.6805
0.5712 | -80.3822
0.57477 | -80.2096
0.57834 | -80.0666
0.58191 | -80.0816
0.58548 | -80.4718
0.58905 | -80.6358
0.59262 | -80.1833
0.59619 | -79.7173
0.59976 | -79.9374
0.60333 | -80.6147
0.6069 | -81.079
0.61047 | -81.3322
0.61404 | -81.1097
0.61761 | -80.8253
0.62118 | -80.7132
0.62475 | -80.6371
0.62832 | -80.5268
0.63189 | -80.7718
0.63546 | -81.0658
0.63903 | -81.3695
0.6426 | -81.1679
0.64617 | -80.6378
0.64974 | -80.6095
0.65331 | -80.7879
0.65688 | -81.1703
0.66045 | -81.9646
0.66402 | -82.477
0.66759 | -81.9848
0.67116 | -81.5425
0.67473 | -81.4839
0.6783 | -81.5833
0.68187 | -81.8626
0.68544 | -81.7286
0.68901 | -81.2992
0.69258 | -81.3004
0.69615 | -81.5316
0.69972 | -81.7788
0.70329 | -82.1759
0.70686 | -82.4056
0.71043 | -82.1728
0.714 | -81.7454
0.71757 | -81.3906
0.72114 | -81.5722
0.72471 | -82.2213
0.72828 | -82.3963
0.73185 | -82.3566
0.73542 | -82.5533
0.73899 | -82.8037
0.74256 | -83.0747
0.74613 | -83.4026
0.7497 | -83.2487
0.75327 | -82.4887
0.75684 | -82.0219
0.76041 | -81.8721
0.76398 | -81.7957
0.76755 | -81.8898
0.77112 | -82.1539
0.77469 | -82.5811
0.77826 | -83.3239
0.78183 | -83.863
0.7854 | -83.3875
0.78897 | -82.5987
0.79254 | -82.5429
0.79611 | -82.851
0.79968 | -83.062
0.80325 | -83.2324
0.80682 | -83.4501
0.81039 | -83.6223
0.81396 | -83.6179
0.81753 | -83.3267
0.8211 | -83.0197
0.82467 | -83.085
0.82824 | -83.1849
0.83181 | -83.6411
0.83538 | -84.6858
0.83895 | -84.9908
0.84252 | -84.5281
0.84609 | -84.6905
0.84966 | -84.6457
0.85323 | -84.1916
0.8568 | -83.8116
0.86037 | -83.7743
0.86394 | -84.1509
0.86751 | -83.9603
0.87108 | -83.5064
0.87465 | -83.914
0.87822 | -84.287
0.88179 | -84.4818
0.88536 | -84.5473
0.88893 | -84.2946
0.8925 | -83.8426
0.89607 | -83.7495
0.89964 | -84.1973
0.90321 | -84.3294
0.90678 | -84.1447
0.91035 | -83.9187
0.91392 | -83.9607
0.91749 | -84.3827
0.92106 | -84.5997
0.92463 | -84.0163
0.9282 | -83.6173
0.93177 | -83.9886
0.93534 | -84.8572
0.93891 | -85.6885
0.94248 | -85.7454
0.94605 | -84.8335
0.94962 | -84.3447
0.95319 | -84.61
0.95676 | -84.7608
0.96033 | -84.4905
0.9639 | -84.2357
0.96747 | -84.2866
0.97104 | -84.8086
0.97461 | -85.5516
0.97818 | -85.7401
0.98175 | -85.7932
0.98532 | -85.6971
0.98889 | -85.2451
0.99246 | -85.0213
0.99603 | -85.4557
0.9996 | -85.7705
1.0032 | -85.6557
1.0067 | -85.8392
1.0103 | -85.9243
1.0139 | -85.7555
1.0174 | -85.0522
1.021 | -84.7102
1.0246 | -85.2052
1.0282 | -85.9613
1.0317 | -86.2769
1.0353 | -85.8921
1.0389 | -85.205
1.0424 | -84.8501
1.046 | -84.9884
1.0496 | -85.3727
1.0531 | -85.7864
1.0567 | -86.1657
1.0603 | -86.4028
1.0639 | -86.4516
1.0674 | -86.4531
1.071 | -86.4786
1.0746 | -86.4665
1.0781 | -86.3989
1.0817 | -86.4221
1.0853 | -86.5984
1.0888 | -86.6748
1.0924 | -86.5308
1.096 | -86.4018
1.0996 | -86.3903
1.1031 | -86.4932
1.1067 | -86.7438
1.1103 | -87.016
1.1138 | -87.053
1.1174 | -86.7571
1.121 | -86.3315
1.1245 | -86.1309
1.1281 | -86.3279
1.1317 | -86.679
1.1353 | -86.7942
1.1388 | -86.7145
1.1424 | -86.6323
1.146 | -86.6405
1.1495 | -86.8745
1.1531 | -87.197
1.1567 | -87.2764
1.1602 | -87.1223
1.1638 | -86.9967
1.1674 | -87.0122
1.171 | -87.0399
1.1745 | -86.8925
1.1781 | -86.6107
1.1817 | -86.4372
1.1852 | -86.4737
1.1888 | -86.627
1.1924 | -86.8748
1.1959 | -87.1987
1.1995 | -87.4332
1.2031 | -87.4815
1.2067 | -87.4941
1.2102 | -87.5534
1.2138 | -87.5888
1.2174 | -87.6565
1.2209 | -87.8265
1.2245 | -88.0597
1.2281 | -88.3175
1.2316 | -88.5079
1.2352 | -88.4837
1.2388 | -88.1494
1.2424 | -87.6469
1.2459 | -87.2714
1.2495 | -87.1821
1.2531 | -87.3358
1.2566 | -87.579
1.2602 | -87.7899
1.2638 | -87.9209
1.2673 | -87.9976
1.2709 | -88.1383
1.2745 | -88.308
1.2781 | -88.3374
1.2816 | -88.3209
1.2852 | -88.4679
1.2888 | -88.6487
1.2923 | -88.5735
1.2959 | -88.2813
1.2995 | -87.9536
1.303 | -87.6889
1.3066 | -87.7376
1.3102 | -88.2515
1.3138 | -88.7813
1.3173 | -88.6533
1.3209 | -88.1122
1.3245 | -87.8687
1.328 | -88.1198
1.3316 | -88.5172
1.3352 | -88.6811
1.3387 | -88.5588
1.3423 | -88.2913
1.3459 | -88.1312
1.3495 | -88.1978
1.353 | -88.4492
1.3566 | -88.7282
1.3602 | -88.7041
1.3637 | -88.3842
1.3673 | -88.2097
1.3709 | -88.4113
1.3744 | -88.9196
1.378 | -89.4874
1.3816 | -89.8758
1.3852 | -90.0316
1.3887 | -89.8704
1.3923 | -89.2442
1.3959 | -88.4851
1.3994 | -88.1591
1.403 | -88.3419
1.4066 | -88.6334
1.4101 | -88.6926
1.4137 | -88.5925
1.4173 | -88.6114
1.4209 | -88.8019
1.4244 | -88.9646
1.428 | -89.0696
1.4316 | -89.2228
1.4351 | -89.2724
1.4387 | -89.0873
1.4423 | -88.9462
1.4458 | -89.0443
1.4494 | -89.1245
1.453 | -88.938
1.4566 | -88.7765
1.4601 | -89.0224
1.4637 | -89.5638
1.4673 | -89.9354
1.4708 | -89.9368
1.4744 | -89.7964
1.478 | -89.7115
1.4815 | -89.6154
1.4851 | -89.3563
1.4887 | -89.0074
1.4923 | -88.7382
1.4958 | -88.5822
1.4994 | -88.5189
1.503 | -88.5714
1.5065 | -88.7523
1.5101 | -88.9874
1.5137 | -89.2652
1.5172 | -89.6561
1.5208 | -90.0386
1.5244 | -90.0806
1.528 | -89.7939
1.5315 | -89.5754
1.5351 | -89.5307
1.5387 | -89.449
1.5422 | -89.4192
1.5458 | -89.5188
1.5494 | -89.7582
1.5529 | -90.0599
1.5565 | -90.2975
1.5601 | -90.4092
1.5637 | -90.3835
1.5672 | -90.1827
1.5708 | -89.8162
1.5744 | -89.4527
1.5779 | -89.2557
1.5815 | -89.2168
1.5851 | -89.2397
1.5886 | -89.3048
1.5922 | -89.4381
1.5958 | -89.6131
1.5994 | -89.7747
1.6029 | -89.8944
1.6065 | -89.9833
1.6101 | -90.0702
1.6136 | -90.1393
1.6172 | -90.1161
1.6208 | -89.9803
1.6243 | -89.8246
1.6279 | -89.7308
1.6315 | -89.7022
1.6351 | -89.7181
1.6386 | -89.7579
1.6422 | -89.8463
1.6458 | -89.9906
1.6493 | -90.1167
1.6529 | -90.1129
1.6565 | -89.9905
1.66 | -89.8836
1.6636 | -89.8921
1.6672 | -90.009
1.6708 | -90.1803
1.6743 | -90.3668
1.6779 | -90.5814
1.6815 | -90.8034
1.685 | -90.9728
1.6886 | -91.0262
1.6922 | -90.9905
1.6957 | -90.8883
1.6993 | -90.7499
1.7029 | -90.6302
1.7065 | -90.5697
1.71 | -90.5319
1.7136 | -90.5018
1.7172 | -90.5373
1.7207 | -90.6702
1.7243 | -90.8396
1.7279 | -90.9628
1.7314 | -90.9698
1.735 | -90.8656
1.7386 | -90.7153
1.7422 | -90.6152
1.7457 | -90.6148
1.7493 | -90.6454
1.7529 | -90.5895
1.7564 | -90.4149
1.76 | -90.2587
1.7636 | -90.2761
1.7671 | -90.479
1.7707 | -90.7743
1.7743 | -91.042
1.7779 | -91.2147
1.7814 | -91.3034
1.785 | -91.2946
1.7886 | -91.1647
1.7921 | -90.9269
1.7957 | -90.6937
1.7993 | -90.5574
1.8028 | -90.5289
1.8064 | -90.5453
1.81 | -90.5567
1.8136 | -90.5701
1.8171 | -90.6252
1.8207 | -90.7402
1.8243 | -90.8756
1.8278 | -90.9314
1.8314 | -90.8473
1.835 | -90.645
1.8385 | -90.4282
1.8421 | -90.2812
1.8457 | -90.2318
1.8493 | -90.3205
1.8528 | -90.5679
1.8564 | -90.9093
1.86 | -91.2132
1.8635 | -91.4085
1.8671 | -91.5509
1.8707 | -91.6869
1.8742 | -91.7804
1.8778 | -91.8301
1.8814 | -91.8388
1.885 | -91.7732
1.8885 | -91.5891
1.8921 | -91.3209
1.8957 | -91.1191
1.8992 | -91.1337
1.9028 | -91.3894
1.9064 | -91.7349
1.9099 | -91.9407
1.9135 | -91.9357
1.9171 | -91.8245
1.9207 | -91.7452
1.9242 | -91.7872
1.9278 | -91.9426
1.9314 | -92.0974
1.9349 | -92.0693
1.9385 | -91.8385
1.9421 | -91.5825
1.9456 | -91.4564
1.9492 | -91.4888
1.9528 | -91.5984
1.9564 | -91.701
1.9599 | -91.7434
1.9635 | -91.6751
1.9671 | -91.5394
1.9706 | -91.4153
1.9742 | -91.3699
1.9778 | -91.4418
1.9813 | -91.6031
1.9849 | -91.7787
1.9885 | -91.8739
1.9921 | -91.8562
1.9956 | -91.8015
1.9992 | -91.7924
2.0028 | -91.8612
2.0063 | -91.9609
2.0099 | -92.0053
2.0135 | -91.971
2.017 | -91.8712
2.0206 | -91.7517
2.0242 | -91.6298
2.0278 | -91.5215
2.0313 | -91.4782
2.0349 | -91.5457
2.0385 | -91.6794
2.042 | -91.7584
2.0456 | -91.7054
2.0492 | -91.6118
2.0527 | -91.5832
2.0563 | -91.6397
2.0599 | -91.7183
2.0635 | -91.7811
2.067 | -91.8351
2.0706 | -91.8743
2.0742 | -91.9001
2.0777 | -91.9256
2.0813 | -91.9914
2.0849 | -92.118
2.0884 | -92.2768
2.092 | -92.4054
2.0956 | -92.4598
2.0992 | -92.4537
2.1027 | -92.3904
2.1063 | -92.2868
2.1099 | -92.1578
2.1134 | -92.0471
2.117 | -92.0095
2.1206 | -92.0751
2.1241 | -92.1797
2.1277 | -92.2352
2.1313 | -92.1783
2.1349 | -92.0269
2.1384 | -91.868
2.142 | -91.7985
2.1456 | -91.8353
2.1491 | -91.9268
2.1527 | -91.9832
2.1563 | -91.9414
2.1598 | -91.8012
2.1634 | -91.6466
2.167 | -91.5783
2.1706 | -91.6572
2.1741 | -91.8767
2.1777 | -92.1796
2.1813 | -92.466
2.1848 | -92.668
2.1884 | -92.7693
2.192 | -92.8016
2.1955 | -92.7949
2.1991 | -92.7804
2.2027 | -92.7761
2.2063 | -92.7655
2.2098 | -92.7282
2.2134 | -92.6847
2.217 | -92.6605
2.2205 | -92.6837
2.2241 | -92.7246
2.2277 | -92.7212
2.2312 | -92.64
2.2348 | -92.516
2.2384 | -92.4075
2.242 | -92.3433
2.2455 | -92.3311
2.2491 | -92.3569
2.2527 | -92.3913
2.2562 | -92.4385
2.2598 | -92.497
2.2634 | -92.5857
2.2669 | -92.7052
2.2705 | -92.8552
2.2741 | -92.9962
2.2777 | -93.0875
2.2812 | -93.0947
2.2848 | -93.0118
2.2884 | -92.8852
2.2919 | -92.7574
2.2955 | -92.6423
2.2991 | -92.535
2.3026 | -92.4486
2.3062 | -92.4283
2.3098 | -92.4866
2.3134 | -92.6113
2.3169 | -92.7692
2.3205 | -92.9313
2.3241 | -93.0699
2.3276 | -93.1811
2.3312 | -93.2579
2.3348 | -93.2796
2.3383 | -93.2432
2.3419 | -93.142
2.3455 | -93.0113
2.3491 | -92.8525
2.3526 | -92.6604
2.3562 | -92.4773
2.3598 | -92.3968
2.3633 | -92.4816
2.3669 | -92.7079
2.3705 | -92.9953
2.374 | -93.2427
2.3776 | -93.3884
2.3812 | -93.4369
2.3848 | -93.3905
2.3883 | -93.3004
2.3919 | -93.2085
2.3955 | -93.1501
2.399 | -93.1544
2.4026 | -93.2097
2.4062 | -93.2796
2.4097 | -93.3284
2.4133 | -93.3524
2.4169 | -93.3391
2.4205 | -93.2721
2.424 | -93.1542
2.4276 | -92.9968
2.4312 | -92.839
2.4347 | -92.7413
2.4383 | -92.7492
2.4419 | -92.8686
2.4454 | -93.0455
2.449 | -93.1795
2.4526 | -93.2059
2.4562 | -93.1374
2.4597 | -93.0522
2.4633 | -93.037
2.4669 | -93.1341
2.4704 | -93.3426
2.474 | -93.5883
2.4776 | -93.7761
2.4811 | -93.8483
2.4847 | -93.819
2.4883 | -93.7569
2.4919 | -93.7133
2.4954 | -93.7048
2.499 | -93.6893
2.5026 | -93.6173
2.5061 | -93.4944
2.5097 | -93.3485
2.5133 | -93.2545
2.5168 | -93.2468
2.5204 | -93.3252
2.524 | -93.4564
2.5276 | -93.6233
2.5311 | -93.8029
2.5347 | -93.9728
2.5383 | -94.1121
2.5418 | -94.1883
2.5454 | -94.1915
2.549 | -94.1367
2.5525 | -94.012
2.5561 | -93.8501
2.5597 | -93.7214
2.5633 | -93.6937
2.5668 | -93.7707
2.5704 | -93.9202
2.574 | -94.0538
2.5775 | -94.1674
2.5811 | -94.2631
2.5847 | -94.3363
2.5882 | -94.3687
2.5918 | -94.3433
2.5954 | -94.2723
2.599 | -94.1907
2.6025 | -94.1341
2.6061 | -94.1302
2.6097 | -94.1628
2.6132 | -94.1894
2.6168 | -94.172
2.6204 | -94.1139
2.6239 | -94.0425
2.6275 | -93.9966
2.6311 | -94.0017
2.6347 | -94.0414
2.6382 | -94.083
2.6418 | -94.0945
2.6454 | -94.061
2.6489 | -94.0032
2.6525 | -93.9405
2.6561 | -93.8933
2.6596 | -93.8667
2.6632 | -93.866
2.6668 | -93.8829
2.6704 | -93.921
2.6739 | -93.9749
2.6775 | -94.0378
2.6811 | -94.0776
2.6846 | -94.0852
2.6882 | -94.0746
2.6918 | -94.0878
2.6953 | -94.1462
2.6989 | -94.2608
2.7025 | -94.4018
2.7061 | -94.5108
2.7096 | -94.5626
2.7132 | -94.5692
2.7168 | -94.5494
2.7203 | -94.5264
2.7239 | -94.5094
2.7275 | -94.5036
2.731 | -94.5007
2.7346 | -94.4965
2.7382 | -94.4843
2.7418 | -94.4672
2.7453 | -94.4426
2.7489 | -94.4237
2.7525 | -94.4371
2.756 | -94.4885
2.7596 | -94.5811
2.7632 | -94.6996
2.7667 | -94.8287
2.7703 | -94.9387
2.7739 | -95.0095
2.7775 | -95.022
2.781 | -94.9834
2.7846 | -94.893
2.7882 | -94.7663
2.7917 | -94.6182
2.7953 | -94.4785
2.7989 | -94.3719
2.8024 | -94.3209
2.806 | -94.3173
2.8096 | -94.3441
2.8132 | -94.388
2.8167 | -94.4641
2.8203 | -94.5875
2.8239 | -94.7604
2.8274 | -94.9443
2.831 | -95.115
2.8346 | -95.2551
2.8381 | -95.349
2.8417 | -95.3818
2.8453 | -95.3623
2.8489 | -95.293
2.8524 | -95.2167
2.856 | -95.174
2.8596 | -95.1769
2.8631 | -95.2314
2.8667 | -95.3227
2.8703 | -95.4338
2.8738 | -95.5519
2.8774 | -95.6523
2.881 | -95.7192
2.8846 | -95.7637
2.8881 | -95.7844
2.8917 | -95.7729
2.8953 | -95.7387
2.8988 | -95.6797
2.9024 | -95.5957
2.906 | -95.4859
2.9095 | -95.3513
2.9131 | -95.2205
2.9167 | -95.1158
2.9203 | -95.0612
2.9238 | -95.0497
2.9274 | -95.0417
2.931 | -95.0026
2.9345 | -94.9296
2.9381 | -94.8482
2.9417 | -94.7975
2.9452 | -94.7959
2.9488 | -94.8578
2.9524 | -94.964
2.956 | -95.0909
2.9595 | -95.1942
2.9631 | -95.2635
2.9667 | -95.2918
2.9702 | -95.2852
2.9738 | -95.247
2.9774 | -95.1869
2.9809 | -95.1068
2.9845 | -95.0155
2.9881 | -94.9283
2.9917 | -94.871
2.9952 | -94.8593
2.9988 | -94.9063
3.0024 | -94.9932
3.0059 | -95.0874
3.0095 | -95.1492
3.0131 | -95.1712
3.0166 | -95.1651
3.0202 | -95.1508
3.0238 | -95.1317
3.0274 | -95.1213
3.0309 | -95.1377
3.0345 | -95.173
3.0381 | -95.2107
3.0416 | -95.2509
3.0452 | -95.2794
3.0488 | -95.2958
3.0523 | -95.3215
3.0559 | -95.3633
3.0595 | -95.4448
3.0631 | -95.5858
3.0666 | -95.7854
3.0702 | -96.0236
3.0738 | -96.2511
3.0773 | -96.4057
3.0809 | -96.4544
3.0845 | -96.3729
3.088 | -96.0898
3.0916 | -95.8863
3.0952 | -95.6928
3.0988 | -95.5326
3.1023 | -95.4279
3.1059 | -95.3857
3.1095 | -95.3945
3.113 | -95.4425
3.1166 | -95.5089
3.1202 | -95.6003
3.1237 | -95.7094
3.1273 | -95.8201
3.1309 | -95.9019
3.1345 | -95.9481
3.138 | -95.9685
3.1416 | -95.9501
3.1452 | -95.8877
3.1487 | -95.7902
3.1523 | -95.6892
3.1559 | -95.6076
3.1594 | -95.5628
3.163 | -95.5635
3.1666 | -95.6073
3.1702 | -95.663
3.1737 | -95.7013
3.1773 | -95.7106
3.1809 | -95.7014
3.1844 | -95.6951
3.188 | -95.6924
3.1916 | -95.7081
3.1951 | -95.7449
3.1987 | -95.7907
3.2023 | -95.8305
3.2059 | -95.8457
3.2094 | -95.8237
3.213 | -95.7817
3.2166 | -95.7276
3.2201 | -95.6694
3.2237 | -95.6329
3.2273 | -95.6159
3.2308 | -95.6352
3.2344 | -95.6985
3.238 | -95.8074
3.2416 | -95.9571
3.2451 | -96.1164
3.2487 | -96.2693
3.2523 | -96.3903
3.2558 | -96.4489
3.2594 | -96.4527
3.263 | -96.421
3.2665 | -96.3605
3.2701 | -96.283
3.2737 | -96.2068
3.2773 | -96.1476
3.2808 | -96.1263
3.2844 | -96.1601
3.288 | -96.2364
3.2915 | -96.3387
3.2951 | -96.4406
3.2987 | -96.5134
3.3022 | -96.5419
3.3058 | -96.5302
3.3094 | -96.4817
3.313 | -96.4037
3.3165 | -96.3106
3.3201 | -96.2201
3.3237 | -96.1483
3.3272 | -96.1135
3.3308 | -96.1214
3.3344 | -96.1599
3.3379 | -96.2178
3.3415 | -96.2696
3.3451 | -96.2983
3.3487 | -96.2999
3.3522 | -96.2643
3.3558 | -96.2128
3.3594 | -96.1649
3.3629 | -96.1338
3.3665 | -96.1207
3.3701 | -96.1262
3.3736 | -96.1451
3.3772 | -96.1749
3.3808 | -96.2057
3.3844 | -96.2367
3.3879 | -96.2686
3.3915 | -96.2835
3.3951 | -96.2997
3.3986 | -96.3084
3.4022 | -96.3119
3.4058 | -96.3244
3.4093 | -96.3361
3.4129 | -96.3534
3.4165 | -96.3792
3.4201 | -96.4114
3.4236 | -96.4495
3.4272 | -96.4927
3.4308 | -96.5355
3.4343 | -96.5612
3.4379 | -96.5695
3.4415 | -96.558
3.445 | -96.5418
3.4486 | -96.5154
3.4522 | -96.493
3.4558 | -96.4857
3.4593 | -96.4957
3.4629 | -96.5177
3.4665 | -96.5539
3.47 | -96.6024
3.4736 | -96.6342
3.4772 | -96.635
3.4807 | -96.6022
3.4843 | -96.5403
3.4879 | -96.4654
3.4915 | -96.3919
3.495 | -96.3235
3.4986 | -96.2551
3.5022 | -96.1898
3.5057 | -96.1314
3.5093 | -96.0852
3.5129 | -96.0683
3.5164 | -96.0919
3.52 | -96.1523
3.5236 | -96.2522
3.5272 | -96.382
3.5307 | -96.5165
3.5343 | -96.6289
3.5379 | -96.7148
3.5414 | -96.7605
3.545 | -96.7623
3.5486 | -96.7506
3.5521 | -96.7464
3.5557 | -96.7625
3.5593 | -96.7978
3.5629 | -96.8425
3.5664 | -96.8755
3.57 | -96.8914
3.5736 | -96.8994
3.5771 | -96.8932
3.5807 | -96.888
3.5843 | -96.8948
3.5878 | -96.9263
3.5914 | -96.9844
3.595 | -97.0775
3.5986 | -97.192
3.6021 | -97.2731
3.6057 | -97.3495
3.6093 | -97.4133
3.6128 | -97.4562
3.6164 | -97.4956
3.62 | -97.5333
3.6235 | -97.5847
3.6271 | -97.6384
3.6307 | -97.6884
3.6343 | -97.7321
3.6378 | -97.7561
3.6414 | -97.7359
3.645 | -97.6709
3.6485 | -97.5595
3.6521 | -97.4008
3.6557 | -97.2208
3.6592 | -97.0375
3.6628 | -96.8668
3.6664 | -96.7288
3.67 | -96.6379
3.6735 | -96.6034
3.6771 | -96.6221
3.6807 | -96.6918
3.6842 | -96.7917
3.6878 | -96.8992
3.6914 | -96.986
3.6949 | -97.0382
3.6985 | -97.0558
3.7021 | -97.0696
3.7057 | -97.0947
3.7092 | -97.1462
3.7128 | -97.2149
3.7164 | -97.2731
3.7199 | -97.3162
3.7235 | -97.3376
3.7271 | -97.3365
3.7306 | -97.3276
3.7342 | -97.3274
3.7378 | -97.3442
3.7414 | -97.3817
3.7449 | -97.4305
3.7485 | -97.4815
3.7521 | -97.5178
3.7556 | -97.5369
3.7592 | -97.5352
3.7628 | -97.504
3.7663 | -97.4496
3.7699 | -97.3873
3.7735 | -97.3297
3.7771 | -97.2693
3.7806 | -97.2063
3.7842 | -97.1512
3.7878 | -97.0994
3.7913 | -97.0628
3.7949 | -97.0504
3.7985 | -97.0668
3.802 | -97.1117
3.8056 | -97.1867
3.8092 | -97.2913
3.8128 | -97.4192
3.8163 | -97.5576
3.8199 | -97.6972
3.8235 | -97.8201
3.827 | -97.931
3.8306 | -97.9984
3.8342 | -98.0257
3.8377 | -98.0201
3.8413 | -97.9788
3.8449 | -97.9028
3.8485 | -97.8242
3.852 | -97.7366
3.8556 | -97.641
3.8592 | -97.5453
3.8627 | -97.4669
3.8663 | -97.3994
3.8699 | -97.3313
3.8734 | -97.2764
3.877 | -97.2432
3.8806 | -97.2331
3.8842 | -97.2581
3.8877 | -97.2964
3.8913 | -97.3553
3.8949 | -97.4164
3.8984 | -97.4658
3.902 | -97.5021
3.9056 | -97.5322
3.9091 | -97.544
3.9127 | -97.5543
3.9163 | -97.5674
3.9199 | -97.5763
3.9234 | -97.5856
3.927 | -97.5941
3.9306 | -97.6021
3.9341 | -97.5867
3.9377 | -97.5487
3.9413 | -97.5059
3.9448 | -97.4712
3.9484 | -97.4608
3.952 | -97.4848
3.9556 | -97.5451
3.9591 | -97.6351
3.9627 | -97.7325
3.9663 | -97.8207
3.9698 | -97.8874
3.9734 | -97.9246
3.977 | -97.947
3.9805 | -97.9629
3.9841 | -97.9699
3.9877 | -97.9795
3.9913 | -97.981
3.9948 | -97.9755
3.9984 | -97.9629
4.002 | -97.9455
4.0055 | -97.9279
4.0091 | -97.9222
4.0127 | -97.9267
4.0162 | -97.9432
4.0198 | -97.9777
4.0234 | -98.0212
4.027 | -98.068
4.0305 | -98.1177
4.0341 | -98.1702
4.0377 | -98.2207
4.0412 | -98.2688
4.0448 | -98.3261
4.0484 | -98.3919
4.0519 | -98.4732
4.0555 | -98.5687
4.0591 | -98.6557
4.0627 | -98.7153
4.0662 | -98.7363
4.0698 | -98.7081
4.0734 | -98.6371
4.0769 | -98.539
4.0805 | -98.4281
4.0841 | -98.3224
4.0876 | -98.2148
4.0912 | -98.1108
4.0948 | -98.0254
4.0984 | -97.9617
4.1019 | -97.9235
4.1055 | -97.8995
4.1091 | -97.8871
4.1126 | -97.8692
4.1162 | -97.8521
4.1198 | -97.855
4.1233 | -97.8551
4.1269 | -97.865
4.1305 | -97.8878
4.1341 | -97.9196
4.1376 | -97.9553
4.1412 | -97.9932
4.1448 | -98.0251
4.1483 | -98.0431
4.1519 | -98.0546
4.1555 | -98.0587
4.159 | -98.0518
4.1626 | -98.0426
4.1662 | -98.0355
4.1698 | -98.0363
4.1733 | -98.038
4.1769 | -98.042
4.1805 | -98.0392
4.184 | -98.031
4.1876 | -98.0186
4.1912 | -98.0034
4.1947 | -97.9867
4.1983 | -97.9703
4.2019 | -97.9666
4.2055 | -97.9875
4.209 | -98.0302
4.2126 | -98.0935
4.2162 | -98.1837
4.2197 | -98.293
4.2233 | -98.4159
4.2269 | -98.5392
4.2304 | -98.6638
4.234 | -98.7796
4.2376 | -98.8667
4.2412 | -98.9269
4.2447 | -98.9489
4.2483 | -98.9304
4.2519 | -98.8756
4.2554 | -98.7948
4.259 | -98.6887
4.2626 | -98.5784
4.2661 | -98.4715
4.2697 | -98.3632
4.2733 | -98.2571
4.2769 | -98.166
4.2804 | -98.0919
4.284 | -98.0333
4.2876 | -97.9919
4.2911 | -97.9698
4.2947 | -97.9767
4.2983 | -97.9965
4.3018 | -98.0271
4.3054 | -98.0691
4.309 | -98.1184
4.3126 | -98.1833
4.3161 | -98.2598
4.3197 | -98.3483
4.3233 | -98.4441
4.3268 | -98.5457
4.3304 | -98.6306
4.334 | -98.6935
4.3375 | -98.738
4.3411 | -98.7584
4.3447 | -98.765
4.3483 | -98.7582
4.3518 | -98.7481
4.3554 | -98.7414
4.359 | -98.7301
4.3625 | -98.7132
4.3661 | -98.7002
4.3697 | -98.6833
4.3732 | -98.6648
4.3768 | -98.6458
4.3804 | -98.6237
4.384 | -98.6027
4.3875 | -98.5875
4.3911 | -98.5751
4.3947 | -98.565
4.3982 | -98.5643
4.4018 | -98.5645
4.4054 | -98.5645
4.4089 | -98.5699
4.4125 | -98.573
4.4161 | -98.5776
4.4197 | -98.5775
4.4232 | -98.5739
4.4268 | -98.5866
4.4304 | -98.6011
4.4339 | -98.6175
4.4375 | -98.6427
4.4411 | -98.6628
4.4446 | -98.6874
4.4482 | -98.7177
4.4518 | -98.7573
4.4554 | -98.8045
4.4589 | -98.8598
4.4625 | -98.9303
4.4661 | -99.0043
4.4696 | -99.0877
4.4732 | -99.1697
4.4768 | -99.2423
4.4803 | -99.3017
4.4839 | -99.3495
4.4875 | -99.3788
4.4911 | -99.3847
4.4946 | -99.3614
4.4982 | -99.3289
4.5018 | -99.2781
4.5053 | -99.2045
4.5089 | -99.1245
4.5125 | -99.0337
4.516 | -98.9486
4.5196 | -98.8756
4.5232 | -98.8128
4.5268 | -98.7652
4.5303 | -98.7302
4.5339 | -98.7125
4.5375 | -98.7024
4.541 | -98.7038
4.5446 | -98.7153
4.5482 | -98.7305
4.5517 | -98.7425
4.5553 | -98.7415
4.5589 | -98.7282
4.5625 | -98.6871
4.566 | -98.6176
4.5696 | -98.5343
4.5732 | -98.4419
4.5767 | -98.3423
4.5803 | -98.245
4.5839 | -98.1589
4.5874 | -98.0953
4.591 | -98.0636
4.5946 | -98.0716
4.5982 | -98.1202
4.6017 | -98.2109
4.6053 | -98.3422
4.6089 | -98.4922
4.6124 | -98.6503
4.616 | -98.8002
4.6196 | -98.9143
4.6231 | -98.9822
4.6267 | -99.0134
4.6303 | -99.0237
4.6339 | -99.0287
4.6374 | -99.0479
4.641 | -99.0711
4.6446 | -99.101
4.6481 | -99.1358
4.6517 | -99.1685
4.6553 | -99.1967
4.6588 | -99.2228
4.6624 | -99.2473
4.666 | -99.261
4.6696 | -99.2745
4.6731 | -99.2831
4.6767 | -99.2865
4.6803 | -99.2976
4.6838 | -99.3086
4.6874 | -99.3287
4.691 | -99.3612
4.6945 | -99.4019
4.6981 | -99.4403
4.7017 | -99.4823
4.7053 | -99.517
4.7088 | -99.5462
4.7124 | -99.5718
4.716 | -99.5817
4.7195 | -99.5864
4.7231 | -99.5845
4.7267 | -99.582
4.7302 | -99.5746
4.7338 | -99.5576
4.7374 | -99.5361
4.741 | -99.5114
4.7445 | -99.4827
4.7481 | -99.461
4.7517 | -99.4445
4.7552 | -99.4328
4.7588 | -99.4324
4.7624 | -99.4436
4.7659 | -99.4697
4.7695 | -99.5025
4.7731 | -99.5338
4.7767 | -99.5652
4.7802 | -99.5963
4.7838 | -99.6151
4.7874 | -99.6243
4.7909 | -99.6335
4.7945 | -99.6332
4.7981 | -99.6286
4.8016 | -99.6209
4.8052 | -99.6144
4.8088 | -99.6063
4.8124 | -99.5882
4.8159 | -99.5677
4.8195 | -99.5371
4.8231 | -99.4984
4.8266 | -99.4601
4.8302 | -99.4202
4.8338 | -99.3839
4.8373 | -99.356
4.8409 | -99.3307
4.8445 | -99.3186
4.8481 | -99.3064
4.8516 | -99.282
4.8552 | -99.2525
4.8588 | -99.2232
4.8623 | -99.1867
4.8659 | -99.1427
4.8695 | -99.1077
4.873 | -99.0787
4.8766 | -99.0549
4.8802 | -99.0372
4.8838 | -99.0349
4.8873 | -99.0441
4.8909 | -99.0529
4.8945 | -99.0664
4.898 | -99.0905
4.9016 | -99.1166
4.9052 | -99.1434
4.9087 | -99.1618
4.9123 | -99.1732
4.9159 | -99.1775
4.9195 | -99.1712
4.923 | -99.1679
4.9266 | -99.1655
4.9302 | -99.1683
4.9337 | -99.1681
4.9373 | -99.1726
4.9409 | -99.1843
4.9444 | -99.1961
4.948 | -99.2157
4.9516 | -99.2426
4.9552 | -99.2832
4.9587 | -99.3384
4.9623 | -99.4025
4.9659 | -99.474
4.9694 | -99.54
4.973 | -99.5837
4.9766 | -99.6033
4.9801 | -99.5876
4.9837 | -99.545
4.9873 | -99.4715
4.9909 | -99.3799
4.9944 | -99.2771
4.998 | -99.1743
5.0016 | -99.0825
5.0051 | -99.0005
5.0087 | -98.9466
5.0123 | -98.9169
5.0158 | -98.9065
5.0194 | -98.9282
5.023 | -98.9781
5.0266 | -99.0549
5.0301 | -99.152
5.0337 | -99.2588
5.0373 | -99.382
5.0408 | -99.5073
5.0444 | -99.6257
5.048 | -99.7355
5.0515 | -99.8294
5.0551 | -99.8995
5.0587 | -99.938
5.0623 | -99.9532
5.0658 | -99.9474
5.0694 | -99.9233
5.073 | -99.8837
5.0765 | -99.8444
5.0801 | -99.803
5.0837 | -99.7707
5.0872 | -99.7528
5.0908 | -99.7541
5.0944 | -99.7725
5.098 | -99.8117
5.1015 | -99.8688
5.1051 | -99.9361
5.1087 | -100.0193
5.1122 | -100.0981
5.1158 | -100.1707
5.1194 | -100.2281
5.1229 | -100.2641
5.1265 | -100.281
5.1301 | -100.28
5.1337 | -100.26
5.1372 | -100.2309
5.1408 | -100.1975
5.1444 | -100.1606
5.1479 | -100.1381
5.1515 | -100.1124
5.1551 | -100.0856
5.1586 | -100.0632
5.1622 | -100.0321
5.1658 | -100.0032
5.1694 | -99.9735
5.1729 | -99.9376
5.1765 | -99.9059
5.1801 | -99.8783
5.1836 | -99.8539
5.1872 | -99.8321
5.1908 | -99.8166
5.1943 | -99.8168
5.1979 | -99.8342
5.2015 | -99.863
5.2051 | -99.9007
5.2086 | -99.9485
5.2122 | -100.0017
5.2158 | -100.0499
5.2193 | -100.0968
5.2229 | -100.1349
5.2265 | -100.1523
5.23 | -100.1572
5.2336 | -100.1449
5.2372 | -100.1138
5.2408 | -100.0649
5.2443 | -100.006
5.2479 | -99.931
5.2515 | -99.8442
5.255 | -99.7536
5.2586 | -99.6681
5.2622 | -99.5976
5.2657 | -99.5478
5.2693 | -99.5157
5.2729 | -99.5134
5.2765 | -99.5244
5.28 | -99.5478
5.2836 | -99.5848
5.2872 | -99.6365
5.2907 | -99.6899
5.2943 | -99.7336
5.2979 | -99.7707
5.3014 | -99.8018
5.305 | -99.8253
5.3086 | -99.843
5.3122 | -99.8547
5.3157 | -99.87
5.3193 | -99.892
5.3229 | -99.9251
5.3264 | -99.9628
5.33 | -100.005
5.3336 | -100.0621
5.3371 | -100.1224
5.3407 | -100.1768
5.3443 | -100.2266
5.3479 | -100.2619
5.3514 | -100.286
5.355 | -100.2874
5.3586 | -100.2853
5.3621 | -100.2769
5.3657 | -100.2509
5.3693 | -100.2071
5.3728 | -100.1594
5.3764 | -100.1163
5.38 | -100.0742
5.3836 | -100.0424
5.3871 | -100.0233
5.3907 | -100.014
5.3943 | -100.0111
5.3978 | -100.0141
5.4014 | -100.0198
5.405 | -100.0283
5.4085 | -100.0412
5.4121 | -100.0494
5.4157 | -100.0547
5.4193 | -100.0596
5.4228 | -100.069
5.4264 | -100.0756
5.43 | -100.0873
5.4335 | -100.0914
5.4371 | -100.0932
5.4407 | -100.0897
5.4442 | -100.0868
5.4478 | -100.0821
5.4514 | -100.0757
5.455 | -100.0748
5.4585 | -100.076
5.4621 | -100.0766
5.4657 | -100.0722
5.4692 | -100.0815
5.4728 | -100.0981
5.4764 | -100.1183
5.4799 | -100.1327
5.4835 | -100.1448
5.4871 | -100.1528
5.4907 | -100.155
5.4942 | -100.1513
5.4978 | -100.1322
5.5014 | -100.1082
5.5049 | -100.0759
5.5085 | -100.0452
5.5121 | -100.0218
5.5156 | -100.0093
5.5192 | -100.0045
5.5228 | -100.0089
5.5264 | -100.0275
5.5299 | -100.0639
5.5335 | -100.1053
5.5371 | -100.1423
5.5406 | -100.1729
5.5442 | -100.1981
5.5478 | -100.2123
5.5513 | -100.2077
5.5549 | -100.1918
5.5585 | -100.1695
5.5621 | -100.1554
5.5656 | -100.1498
5.5692 | -100.1598
5.5728 | -100.1902
5.5763 | -100.2458
5.5799 | -100.3218
5.5835 | -100.4072
5.587 | -100.5048
5.5906 | -100.6
5.5942 | -100.6826
5.5978 | -100.7491
5.6013 | -100.7932
5.6049 | -100.8146
5.6085 | -100.8132
5.612 | -100.7928
5.6156 | -100.7583
5.6192 | -100.7184
5.6227 | -100.6802
5.6263 | -100.6437
5.6299 | -100.6122
5.6335 | -100.5907
5.637 | -100.5789
5.6406 | -100.58
5.6442 | -100.5901
5.6477 | -100.6059
5.6513 | -100.6184
5.6549 | -100.6332
5.6584 | -100.6378
5.662 | -100.6012
5.6656 | -100.5761
5.6692 | -100.5361
5.6727 | -100.4829
5.6763 | -100.4212
5.6799 | -100.3562
5.6834 | -100.2904
5.687 | -100.2235
5.6906 | -100.1594
5.6941 | -100.1028
5.6977 | -100.0577
5.7013 | -100.0294
5.7049 | -100.0123
5.7084 | -100.014
5.712 | -100.0338
5.7156 | -100.0734
5.7191 | -100.1262
5.7227 | -100.1926
5.7263 | -100.2649
5.7298 | -100.3384
5.7334 | -100.4107
5.737 | -100.4747
5.7406 | -100.5195
5.7441 | -100.5522
5.7477 | -100.5699
5.7513 | -100.5715
5.7548 | -100.5618
5.7584 | -100.5444
5.762 | -100.5165
5.7655 | -100.4839
5.7691 | -100.4489
5.7727 | -100.416
5.7763 | -100.3915
5.7798 | -100.369
5.7834 | -100.3477
5.787 | -100.3307
5.7905 | -100.3147
5.7941 | -100.302
5.7977 | -100.2884
5.8012 | -100.2742
5.8048 | -100.2671
5.8084 | -100.2628
5.812 | -100.2549
5.8155 | -100.2384
5.8191 | -100.2143
5.8227 | -100.1818
5.8262 | -100.1388
5.8298 | -100.0945
5.8334 | -100.0403
5.8369 | -99.9739
5.8405 | -99.9012
5.8441 | -99.8301
5.8476 | -99.7554
5.8512 | -99.6836
5.8548 | -99.6172
5.8584 | -99.5585
5.8619 | -99.5057
5.8655 | -99.4595
5.8691 | -99.4214
5.8726 | -99.3895
5.8762 | -99.3655
5.8798 | -99.351
5.8834 | -99.3502
5.8869 | -99.3596
5.8905 | -99.3754
5.8941 | -99.3897
5.8976 | -99.4138
5.9012 | -99.4443
5.9048 | -99.4786
5.9083 | -99.505
5.9119 | -99.5193
5.9155 | -99.5281
5.9191 | -99.5242
5.9226 | -99.512
5.9262 | -99.4927
5.9298 | -99.467
5.9333 | -99.4342
5.9369 | -99.4036
5.9405 | -99.3739
5.944 | -99.3439
5.9476 | -99.3139
5.9512 | -99.287
5.9547 | -99.267
5.9583 | -99.2546
5.9619 | -99.2462
5.9655 | -99.242
5.969 | -99.2472
5.9726 | -99.253
5.9762 | -99.2652
5.9797 | -99.2932
5.9833 | -99.3299
5.9869 | -99.3681
5.9905 | -99.4101
5.994 | -99.4548
5.9976 | -99.5046
6.0012 | -99.5598
6.0047 | -99.6178
6.0083 | -99.6766
6.0119 | -99.7381
6.0154 | -99.8096
6.019 | -99.886
6.0226 | -99.9637
6.0262 | -100.047
6.0297 | -100.1322
6.0333 | -100.2169
6.0369 | -100.2937
6.0404 | -100.366
6.044 | -100.4273
6.0476 | -100.48
6.0511 | -100.5252
6.0547 | -100.5618
6.0583 | -100.5923
6.0618 | -100.6305
6.0654 | -100.6654
6.069 | -100.6902
6.0726 | -100.7214
6.0761 | -100.7562
6.0797 | -100.7932
6.0833 | -100.8256
6.0868 | -100.8437
6.0904 | -100.8517
6.094 | -100.8494
6.0976 | -100.8383
6.1011 | -100.8239
6.1047 | -100.8005
6.1083 | -100.7693
6.1118 | -100.7417
6.1154 | -100.7123
6.119 | -100.6846
6.1225 | -100.6618
6.1261 | -100.6471
6.1297 | -100.6365
6.1333 | -100.6245
6.1368 | -100.615
6.1404 | -100.6076
6.144 | -100.6028
6.1475 | -100.6002
6.1511 | -100.5969
6.1547 | -100.598
6.1582 | -100.5987
6.1618 | -100.6019
6.1654 | -100.6084
6.1689 | -100.6162
6.1725 | -100.636
6.1761 | -100.6664
6.1797 | -100.7045
6.1832 | -100.7392
6.1868 | -100.7712
6.1904 | -100.8049
6.1939 | -100.8395
6.1975 | -100.8738
6.2011 | -100.9088
6.2046 | -100.9362
6.2082 | -100.9657
6.2118 | -100.9888
6.2154 | -101.0087
6.2189 | -101.0287
6.2225 | -101.0463
6.2261 | -101.0703
6.2296 | -101.0896
6.2332 | -101.1123
6.2368 | -101.1269
6.2404 | -101.1345
6.2439 | -101.1387
6.2475 | -101.1422
6.2511 | -101.1389
6.2546 | -101.1223
6.2582 | -101.1045
6.2618 | -101.0824
6.2653 | -101.0529
6.2689 | -101.0211
6.2725 | -100.9911
6.276 | -100.9562
6.2796 | -100.9178
6.2832 | -100.879
6.2868 | -100.8409
6.2903 | -100.8014
6.2939 | -100.7607
6.2975 | -100.7285
6.301 | -100.7061
6.3046 | -100.6855
6.3082 | -100.6779
6.3117 | -100.6817
6.3153 | -100.6914
6.3189 | -100.7007
6.3225 | -100.7119
6.326 | -100.733
6.3296 | -100.7478
6.3332 | -100.7663
6.3367 | -100.7838
6.3403 | -100.7966
6.3439 | -100.8087
6.3475 | -100.8185
6.351 | -100.8231
6.3546 | -100.825
6.3582 | -100.8268
6.3617 | -100.8215
6.3653 | -100.8115
6.3689 | -100.7925
6.3724 | -100.7754
6.376 | -100.7649
6.3796 | -100.7618
6.3831 | -100.7632
6.3867 | -100.7718
6.3903 | -100.7917
6.3939 | -100.8123
6.3974 | -100.8339
6.401 | -100.8667
6.4046 | -100.9004
6.4081 | -100.9312
6.4117 | -100.9681
6.4153 | -101.0056
6.4188 | -101.0428
6.4224 | -101.0742
6.426 | -101.1065
6.4296 | -101.1321
6.4331 | -101.1566
6.4367 | -101.1743
6.4403 | -101.1836
6.4438 | -101.1892
6.4474 | -101.1888
6.451 | -101.1833
6.4546 | -101.1829
6.4581 | -101.187
6.4617 | -101.1865
6.4653 | -101.194
6.4688 | -101.2078
6.4724 | -101.2216
6.476 | -101.2371
6.4795 | -101.2545
6.4831 | -101.2787
6.4867 | -101.2971
6.4902 | -101.3167
6.4938 | -101.336
6.4974 | -101.3489
6.501 | -101.359
6.5045 | -101.3686
6.5081 | -101.3744
6.5117 | -101.3822
6.5152 | -101.3871
6.5188 | -101.3915
6.5224 | -101.3952
6.5259 | -101.3984
6.5295 | -101.4026
6.5331 | -101.4157
6.5367 | -101.431
6.5402 | -101.4562
6.5438 | -101.4854
6.5474 | -101.5105
6.5509 | -101.5411
6.5545 | -101.5767
6.5581 | -101.6188
6.5616 | -101.6518
6.5652 | -101.671
6.5688 | -101.6909
6.5724 | -101.707
6.5759 | -101.7235
6.5795 | -101.7289
6.5831 | -101.73
6.5866 | -101.7263
6.5902 | -101.7172
6.5938 | -101.7042
6.5973 | -101.6918
6.6009 | -101.6716
6.6045 | -101.6522
6.6081 | -101.6389
6.6116 | -101.6299
6.6152 | -101.6247
6.6188 | -101.6271
6.6223 | -101.6334
6.6259 | -101.6387
6.6295 | -101.6435
6.633 | -101.6521
6.6366 | -101.6603
6.6402 | -101.6732
6.6438 | -101.6911
6.6473 | -101.7034
6.6509 | -101.7207
6.6545 | -101.7408
6.658 | -101.7611
6.6616 | -101.7801
6.6652 | -101.7919
6.6687 | -101.8036
6.6723 | -101.8105
6.6759 | -101.822
6.6795 | -101.8401
6.683 | -101.8571
6.6866 | -101.8774
6.6902 | -101.8955
6.6937 | -101.914
6.6973 | -101.935
6.7009 | -101.9603
6.7044 | -101.9771
6.708 | -101.9908
6.7116 | -102.0021
6.7152 | -102.0125
6.7187 | -102.0186
6.7223 | -102.0148
6.7259 | -102.0078
6.7294 | -101.9958
6.733 | -101.9794
6.7366 | -101.9541
6.7401 | -101.9229
6.7437 | -101.8829
6.7473 | -101.8428
6.7509 | -101.7966
6.7544 | -101.7435
6.758 | -101.6953
6.7616 | -101.6544
6.7651 | -101.6137
6.7687 | -101.5756
6.7723 | -101.5419
6.7758 | -101.5133
6.7794 | -101.4937
6.783 | -101.4841
6.7866 | -101.4816
6.7901 | -101.4898
6.7937 | -101.5057
6.7973 | -101.5297
6.8008 | -101.551
6.8044 | -101.5753
6.808 | -101.604
6.8115 | -101.6289
6.8151 | -101.643
6.8187 | -101.6452
6.8223 | -101.6435
6.8258 | -101.6415
6.8294 | -101.6312
6.833 | -101.6116
6.8365 | -101.592
6.8401 | -101.575
6.8437 | -101.5592
6.8472 | -101.5465
6.8508 | -101.5334
6.8544 | -101.5257
6.858 | -101.5234
6.8615 | -101.5192
6.8651 | -101.5131
6.8687 | -101.5175
6.8722 | -101.5252
6.8758 | -101.5344
6.8794 | -101.5464
6.8829 | -101.5534
6.8865 | -101.5589
6.8901 | -101.5702
6.8937 | -101.5742
6.8972 | -101.5732
6.9008 | -101.5681
6.9044 | -101.5618
6.9079 | -101.5542
6.9115 | -101.5412
6.9151 | -101.5279
6.9186 | -101.5201
6.9222 | -101.5153
6.9258 | -101.5084
6.9294 | -101.5001
6.9329 | -101.4948
6.9365 | -101.4944
6.9401 | -101.4866
6.9436 | -101.4819
6.9472 | -101.4774
6.9508 | -101.4783
6.9543 | -101.4838
6.9579 | -101.4888
6.9615 | -101.4982
6.9651 | -101.5148
6.9686 | -101.542
6.9722 | -101.5697
6.9758 | -101.5988
6.9793 | -101.6333
6.9829 | -101.6731
6.9865 | -101.7182
6.99 | -101.7614
6.9936 | -101.8006
6.9972 | -101.8469
7.0008 | -101.8899
7.0043 | -101.9256
7.0079 | -101.9452
7.0115 | -101.9566
7.015 | -101.96
7.0186 | -101.9454
7.0222 | -101.9147
7.0257 | -101.8728
7.0293 | -101.8191
7.0329 | -101.7541
7.0365 | -101.6834
7.04 | -101.6145
7.0436 | -101.5559
7.0472 | -101.4993
7.0507 | -101.4499
7.0543 | -101.4071
7.0579 | -101.3791
7.0614 | -101.3651
7.065 | -101.3603
7.0686 | -101.3659
7.0722 | -101.3789
7.0757 | -101.4063
7.0793 | -101.4462
7.0829 | -101.4891
7.0864 | -101.5343
7.09 | -101.5861
7.0936 | -101.6335
7.0971 | -101.6799
7.1007 | -101.7237
7.1043 | -101.7662
7.1079 | -101.8071
7.1114 | -101.8501
7.115 | -101.88
7.1186 | -101.9004
7.1221 | -101.9233
7.1257 | -101.9421
7.1293 | -101.9552
7.1328 | -101.9711
7.1364 | -101.9939
7.14 | -102.0201
7.1436 | -102.0466
7.1471 | -102.0746
7.1507 | -102.1096
7.1543 | -102.152
7.1578 | -102.1997
7.1614 | -102.2431
7.165 | -102.2802
7.1685 | -102.3149
7.1721 | -102.3415
7.1757 | -102.3575
7.1793 | -102.362
7.1828 | -102.3527
7.1864 | -102.3373
7.19 | -102.3083
7.1935 | -102.2661
7.1971 | -102.2234
7.2007 | -102.1769
7.2042 | -102.1255
7.2078 | -102.0951
7.2114 | -102.0589
7.215 | -102.0307
7.2185 | -102.0128
7.2221 | -102.0087
7.2257 | -102.0159
7.2292 | -102.0306
7.2328 | -102.056
7.2364 | -102.092
7.2399 | -102.1344
7.2435 | -102.1781
7.2471 | -102.2191
7.2507 | -102.2621
7.2542 | -102.3048
7.2578 | -102.3368
7.2614 | -102.3599
7.2649 | -102.3743
7.2685 | -102.3822
7.2721 | -102.3836
7.2756 | -102.3748
7.2792 | -102.352
7.2828 | -102.3234
7.2864 | -102.2895
7.2899 | -102.254
7.2935 | -102.2102
7.2971 | -102.1652
7.3006 | -102.119
7.3042 | -102.0775
7.3078 | -102.0382
7.3113 | -102.003
7.3149 | -101.9756
7.3185 | -101.9549
7.3221 | -101.9424
7.3256 | -101.9368
7.3292 | -101.9358
7.3328 | -101.9386
7.3363 | -101.9442
7.3399 | -101.9487
7.3435 | -101.9505
7.347 | -101.9528
7.3506 | -101.9532
7.3542 | -101.9589
7.3578 | -101.9663
7.3613 | -101.9682
7.3649 | -101.9693
7.3685 | -101.9739
7.372 | -101.9758
7.3756 | -101.9724
7.3792 | -101.9654
7.3827 | -101.959
7.3863 | -101.9514
7.3899 | -101.9481
7.3935 | -101.9491
7.397 | -101.9506
7.4006 | -101.9609
7.4042 | -101.9811
7.4077 | -102.0088
7.4113 | -102.0443
7.4149 | -102.0786
7.4184 | -102.1164
7.422 | -102.1625
7.4256 | -102.2105
7.4292 | -102.255
7.4327 | -102.2961
7.4363 | -102.3354
7.4399 | -102.3757
7.4434 | -102.4104
7.447 | -102.4437
7.4506 | -102.4761
7.4541 | -102.5027
7.4577 | -102.5196
7.4613 | -102.5334
7.4649 | -102.5416
7.4684 | -102.5446
7.472 | -102.5467
7.4756 | -102.5495
7.4791 | -102.5492
7.4827 | -102.5442
7.4863 | -102.5388
7.4898 | -102.5329
7.4934 | -102.5197
7.497 | -102.4994
7.5006 | -102.476
7.5041 | -102.4496
7.5077 | -102.4228
7.5113 | -102.3957
7.5148 | -102.3657
7.5184 | -102.3299
7.522 | -102.2948
7.5255 | -102.2558
7.5291 | -102.219
7.5327 | -102.1777
7.5363 | -102.138
7.5398 | -102.1042
7.5434 | -102.0735
7.547 | -102.0477
7.5505 | -102.0275
7.5541 | -102.0191
7.5577 | -102.0159
7.5612 | -102.018
7.5648 | -102.0264
7.5684 | -102.0396
7.572 | -102.0605
7.5755 | -102.0829
7.5791 | -102.1072
7.5827 | -102.1308
7.5862 | -102.1507
7.5898 | -102.1717
7.5934 | -102.1977
7.5969 | -102.2201
7.6005 | -102.2379
7.6041 | -102.2521
7.6077 | -102.2644
7.6112 | -102.2773
7.6148 | -102.2893
7.6184 | -102.3015
7.6219 | -102.3154
7.6255 | -102.3376
7.6291 | -102.3649
7.6326 | -102.3888
7.6362 | -102.4253
7.6398 | -102.4607
7.6434 | -102.4998
7.6469 | -102.5394
7.6505 | -102.5786
7.6541 | -102.6131
7.6576 | -102.6413
7.6612 | -102.6601
7.6648 | -102.6686
7.6683 | -102.6752
7.6719 | -102.6791
7.6755 | -102.6815
7.6791 | -102.6733
7.6826 | -102.6556
7.6862 | -102.6435
7.6898 | -102.6352
7.6933 | -102.629
7.6969 | -102.6267
7.7005 | -102.628
7.704 | -102.6255
7.7076 | -102.6274
7.7112 | -102.6332
7.7148 | -102.6427
7.7183 | -102.6454
7.7219 | -102.6398
7.7255 | -102.6216
7.729 | -102.6112
7.7326 | -102.5984
7.7362 | -102.5835
7.7397 | -102.5674
7.7433 | -102.547
7.7469 | -102.5238
7.7505 | -102.4965
7.754 | -102.4709
7.7576 | -102.4461
7.7612 | -102.4252
7.7647 | -102.4133
7.7683 | -102.4048
7.7719 | -102.3996
7.7754 | -102.3949
7.779 | -102.3972
7.7826 | -102.3986
7.7862 | -102.4018
7.7897 | -102.4032
7.7933 | -102.4101
7.7969 | -102.4173
7.8004 | -102.4186
7.804 | -102.416
7.8076 | -102.4129
7.8111 | -102.4098
7.8147 | -102.4019
7.8183 | -102.3974
7.8219 | -102.3995
7.8254 | -102.4068
7.829 | -102.4196
7.8326 | -102.4401
7.8361 | -102.4665
7.8397 | -102.4971
7.8433 | -102.5295
7.8468 | -102.5701
7.8504 | -102.617
7.854 | -102.6674
7.8576 | -102.7165
7.8611 | -102.7738
7.8647 | -102.8349
7.8683 | -102.8858
7.8718 | -102.9333
7.8754 | -102.9838
7.879 | -103.0291
7.8825 | -103.0694
7.8861 | -103.1131
7.8897 | -103.152
7.8933 | -103.1942
7.8968 | -103.2247
7.9004 | -103.2448
7.904 | -103.2559
7.9075 | -103.2618
7.9111 | -103.2557
7.9147 | -103.2374
7.9182 | -103.2135
7.9218 | -103.1849
7.9254 | -103.1507
7.929 | -103.1126
7.9325 | -103.0748
7.9361 | -103.0366
7.9397 | -102.9974
7.9432 | -102.9598
7.9468 | -102.9257
7.9504 | -102.898
7.9539 | -102.8766
7.9575 | -102.8608
7.9611 | -102.8453
7.9647 | -102.8296
7.9682 | -102.8225
7.9718 | -102.818
7.9754 | -102.8167
7.9789 | -102.8135
7.9825 | -102.8121
7.9861 | -102.8194
7.9896 | -102.8285
7.9932 | -102.8305
7.9968 | -102.8314
8.0004 | -102.833
8.0039 | -102.8337
8.0075 | -102.8364
8.0111 | -102.8347
8.0146 | -102.8344
8.0182 | -102.8365
8.0218 | -102.8383
8.0253 | -102.8395
8.0289 | -102.841
8.0325 | -102.8449
8.0361 | -102.8493
8.0396 | -102.858
8.0432 | -102.8616
8.0468 | -102.8665
8.0503 | -102.8776
8.0539 | -102.8907
8.0575 | -102.9037
8.061 | -102.9097
8.0646 | -102.9209
8.0682 | -102.9299
8.0718 | -102.9365
8.0753 | -102.9423
8.0789 | -102.9483
8.0825 | -102.9523
8.086 | -102.9496
8.0896 | -102.9471
8.0932 | -102.9466
8.0967 | -102.9484
8.1003 | -102.95
8.1039 | -102.9512
8.1075 | -102.9551
8.111 | -102.9637
8.1146 | -102.9736
8.1182 | -102.982
8.1217 | -102.986
8.1253 | -102.9875
8.1289 | -102.986
8.1324 | -102.9839
8.136 | -102.976
8.1396 | -102.9632
8.1432 | -102.9435
8.1467 | -102.9255
8.1503 | -102.9035
8.1539 | -102.8792
8.1574 | -102.8551
8.161 | -102.8356
8.1646 | -102.8161
8.1681 | -102.7955
8.1717 | -102.7736
8.1753 | -102.7557
8.1789 | -102.7484
8.1824 | -102.7412
8.186 | -102.7393
8.1896 | -102.7426
8.1931 | -102.7542
8.1967 | -102.7721
8.2003 | -102.7944
8.2038 | -102.8228
8.2074 | -102.8578
8.211 | -102.8966
8.2146 | -102.9394
8.2181 | -102.9857
8.2217 | -103.0319
8.2253 | -103.0783
8.2288 | -103.128
8.2324 | -103.1746
8.236 | -103.2194
8.2395 | -103.2484
8.2431 | -103.2824
8.2467 | -103.3082
8.2503 | -103.3263
8.2538 | -103.3405
8.2574 | -103.3469
8.261 | -103.3446
8.2645 | -103.3402
8.2681 | -103.3216
8.2717 | -103.2942
8.2752 | -103.2633
8.2788 | -103.2301
8.2824 | -103.1947
8.286 | -103.1571
8.2895 | -103.1219
8.2931 | -103.0898
8.2967 | -103.0567
8.3002 | -103.0288
8.3038 | -103.0053
8.3074 | -102.9866
8.3109 | -102.9701
8.3145 | -102.9566
8.3181 | -102.9542
8.3217 | -102.9563
8.3252 | -102.9614
8.3288 | -102.971
8.3324 | -102.9851
8.3359 | -103.0056
8.3395 | -103.0245
8.3431 | -103.0496
8.3466 | -103.0746
8.3502 | -103.0991
8.3538 | -103.1279
8.3574 | -103.1542
8.3609 | -103.1788
8.3645 | -103.2061
8.3681 | -103.2337
8.3716 | -103.2652
8.3752 | -103.2918
8.3788 | -103.3091
8.3823 | -103.3324
8.3859 | -103.3608
8.3895 | -103.3853
8.3931 | -103.4098
8.3966 | -103.4265
8.4002 | -103.4339
8.4038 | -103.4352
8.4073 | -103.4301
8.4109 | -103.4171
8.4145 | -103.3974
8.418 | -103.3697
8.4216 | -103.3368
8.4252 | -103.296
8.4288 | -103.2475
8.4323 | -103.1974
8.4359 | -103.1504
8.4395 | -103.1025
8.443 | -103.0527
8.4466 | -103.0084
8.4502 | -102.9694
8.4537 | -102.9373
8.4573 | -102.9126
8.4609 | -102.8877
8.4645 | -102.875
8.468 | -102.8709
8.4716 | -102.8779
8.4752 | -102.8908
8.4787 | -102.911
8.4823 | -102.9385
8.4859 | -102.9725
8.4894 | -103.0068
8.493 | -103.0402
8.4966 | -103.0731
8.5002 | -103.1102
8.5037 | -103.1461
8.5073 | -103.1701
8.5109 | -103.1858
8.5144 | -103.1998
8.518 | -103.2116
8.5216 | -103.2268
8.5251 | -103.2384
8.5287 | -103.2449
8.5323 | -103.2535
8.5359 | -103.2613
8.5394 | -103.2692
8.543 | -103.2808
8.5466 | -103.2943
8.5501 | -103.3091
8.5537 | -103.3271
8.5573 | -103.3494
8.5608 | -103.3709
8.5644 | -103.3906
8.568 | -103.4081
8.5716 | -103.431
8.5751 | -103.4509
8.5787 | -103.4631
8.5823 | -103.475
8.5858 | -103.486
8.5894 | -103.4891
8.593 | -103.486
8.5965 | -103.4793
8.6001 | -103.4695
8.6037 | -103.4569
8.6073 | -103.4398
8.6108 | -103.4193
8.6144 | -103.3996
8.618 | -103.3826
8.6215 | -103.3677
8.6251 | -103.3533
8.6287 | -103.3399
8.6322 | -103.3306
8.6358 | -103.3245
8.6394 | -103.3188
8.643 | -103.3148
8.6465 | -103.3088
8.6501 | -103.3002
8.6537 | -103.2925
8.6572 | -103.2877
8.6608 | -103.2854
8.6644 | -103.2816
8.6679 | -103.2798
8.6715 | -103.2837
8.6751 | -103.2875
8.6787 | -103.2933
8.6822 | -103.3012
8.6858 | -103.3121
8.6894 | -103.3322
8.6929 | -103.3563
8.6965 | -103.3795
8.7001 | -103.4099
8.7036 | -103.444
8.7072 | -103.4811
8.7108 | -103.5205
8.7144 | -103.5607
8.7179 | -103.5992
8.7215 | -103.6315
8.7251 | -103.6636
8.7286 | -103.6921
8.7322 | -103.7192
8.7358 | -103.7452
8.7393 | -103.7674
8.7429 | -103.784
8.7465 | -103.792
8.7501 | -103.7926
8.7536 | -103.7713
8.7572 | -103.7634
8.7608 | -103.751
8.7643 | -103.7339
8.7679 | -103.716
8.7715 | -103.6946
8.775 | -103.6707
8.7786 | -103.6484
8.7822 | -103.6199
8.7858 | -103.589
8.7893 | -103.564
8.7929 | -103.5417
8.7965 | -103.5161
8.8 | -103.4852
8.8036 | -103.4612
8.8072 | -103.4366
8.8107 | -103.4103
8.8143 | -103.3866
8.8179 | -103.3629
8.8215 | -103.3416
8.825 | -103.3174
8.8286 | -103.2984
8.8322 | -103.2854
8.8357 | -103.2708
8.8393 | -103.2592
8.8429 | -103.258
8.8464 | -103.2611
8.85 | -103.2716
8.8536 | -103.2857
8.8572 | -103.3016
8.8607 | -103.3229
8.8643 | -103.3445
8.8679 | -103.3637
8.8714 | -103.3833
8.875 | -103.4002
8.8786 | -103.4177
8.8821 | -103.4344
8.8857 | -103.4431
8.8893 | -103.4428
8.8929 | -103.4341
8.8964 | -103.4157
8.9 | -103.3886
8.9036 | -103.3554
8.9071 | -103.3174
8.9107 | -103.2771
8.9143 | -103.2355
8.9178 | -103.1854
8.9214 | -103.1312
8.925 | -103.0806
8.9286 | -103.0352
8.9321 | -102.9954
8.9357 | -102.9528
8.9393 | -102.9102
8.9428 | -102.8717
8.9464 | -102.8423
8.95 | -102.8186
8.9535 | -102.8002
8.9571 | -102.7892
8.9607 | -102.7806
8.9643 | -102.7756
8.9678 | -102.7766
8.9714 | -102.7842
8.975 | -102.7917
8.9785 | -102.8033
8.9821 | -102.8193
8.9857 | -102.8376
8.9892 | -102.8624
8.9928 | -102.8919
8.9964 | -102.924
9 | -102.9626
9.0035 | -103.0044
9.0071 | -103.0517
9.0107 | -103.1029
9.0142 | -103.1551
9.0178 | -103.2099
9.0214 | -103.2705
9.0249 | -103.3315
9.0285 | -103.3895
9.0321 | -103.4475
9.0357 | -103.5025
9.0392 | -103.551
9.0428 | -103.5918
9.0464 | -103.6258
9.0499 | -103.6564
9.0535 | -103.6805
9.0571 | -103.6923
9.0606 | -103.6952
9.0642 | -103.6904
9.0678 | -103.6814
9.0714 | -103.6714
9.0749 | -103.6507
9.0785 | -103.6182
9.0821 | -103.5869
9.0856 | -103.5567
9.0892 | -103.528
9.0928 | -103.4986
9.0963 | -103.4724
9.0999 | -103.4569
9.1035 | -103.4448
9.1071 | -103.4358
9.1106 | -103.4338
9.1142 | -103.4336
9.1178 | -103.4364
9.1213 | -103.4411
9.1249 | -103.4493
9.1285 | -103.4657
9.132 | -103.4832
9.1356 | -103.5056
9.1392 | -103.5243
9.1428 | -103.5429
9.1463 | -103.5655
9.1499 | -103.5916
9.1535 | -103.6105
9.157 | -103.629
9.1606 | -103.6515
9.1642 | -103.669
9.1677 | -103.6844
9.1713 | -103.7029
9.1749 | -103.7167
9.1785 | -103.7292
9.182 | -103.7414
9.1856 | -103.7467
9.1892 | -103.7524
9.1927 | -103.7612
9.1963 | -103.7713
9.1999 | -103.7824
9.2034 | -103.7914
9.207 | -103.7988
9.2106 | -103.8085
9.2142 | -103.8139
9.2177 | -103.8181
9.2213 | -103.8238
9.2249 | -103.8315
9.2284 | -103.84
9.232 | -103.8414
9.2356 | -103.8415
9.2391 | -103.8424
9.2427 | -103.8356
9.2463 | -103.8281
9.2499 | -103.8201
9.2534 | -103.8088
9.257 | -103.7939
9.2606 | -103.7774
9.2641 | -103.7627
9.2677 | -103.753
9.2713 | -103.7418
9.2748 | -103.7361
9.2784 | -103.729
9.282 | -103.7192
9.2856 | -103.71
9.2891 | -103.7097
9.2927 | -103.7072
9.2963 | -103.7003
9.2998 | -103.6959
9.3034 | -103.6927
9.307 | -103.689
9.3105 | -103.6829
9.3141 | -103.6768
9.3177 | -103.6695
9.3213 | -103.6623
9.3248 | -103.6544
9.3284 | -103.6438
9.332 | -103.6304
9.3355 | -103.6172
9.3391 | -103.6068
9.3427 | -103.5969
9.3462 | -103.5805
9.3498 | -103.5617
9.3534 | -103.5428
9.357 | -103.5219
9.3605 | -103.5016
9.3641 | -103.4835
9.3677 | -103.4642
9.3712 | -103.4519
9.3748 | -103.4407
9.3784 | -103.4325
9.3819 | -103.4281
9.3855 | -103.4247
9.3891 | -103.4262
9.3927 | -103.4318
9.3962 | -103.4396
9.3998 | -103.4497
9.4034 | -103.4633
9.4069 | -103.4781
9.4105 | -103.4981
9.4141 | -103.5214
9.4176 | -103.5475
9.4212 | -103.5761
9.4248 | -103.6048
9.4284 | -103.6304
9.4319 | -103.6634
9.4355 | -103.7002
9.4391 | -103.7347
9.4426 | -103.7687
9.4462 | -103.8033
9.4498 | -103.841
9.4533 | -103.8749
9.4569 | -103.9101
9.4605 | -103.9432
9.4641 | -103.9703
9.4676 | -103.991
9.4712 | -104.009
9.4748 | -104.0203
9.4783 | -104.0263
9.4819 | -104.0321
9.4855 | -104.0356
9.489 | -104.0299
9.4926 | -104.0181
9.4962 | -104.0038
9.4998 | -103.9851
9.5033 | -103.9626
9.5069 | -103.9424
9.5105 | -103.9211
9.514 | -103.9048
9.5176 | -103.8966
9.5212 | -103.8884
9.5247 | -103.878
9.5283 | -103.8713
9.5319 | -103.8722
9.5355 | -103.8767
9.539 | -103.8867
9.5426 | -103.9006
9.5462 | -103.9219
9.5497 | -103.9427
9.5533 | -103.9637
9.5569 | -103.989
9.5604 | -104.0162
9.564 | -104.0454
9.5676 | -104.0721
9.5712 | -104.0948
9.5747 | -104.1133
9.5783 | -104.1306
9.5819 | -104.1474
9.5854 | -104.1581
9.589 | -104.1658
9.5926 | -104.1747
9.5961 | -104.1712
9.5997 | -104.1662
9.6033 | -104.1581
9.6069 | -104.1473
9.6104 | -104.1311
9.614 | -104.1098
9.6176 | -104.0833
9.6211 | -104.0527
9.6247 | -104.0223
9.6283 | -103.9946
9.6318 | -103.9627
9.6354 | -103.9256
9.639 | -103.8895
9.6426 | -103.8551
9.6461 | -103.8223
9.6497 | -103.7924
9.6533 | -103.7672
9.6568 | -103.7459
9.6604 | -103.7307
9.664 | -103.7219
9.6675 | -103.7156
9.6711 | -103.7104
9.6747 | -103.7117
9.6783 | -103.7153
9.6818 | -103.7209
9.6854 | -103.7292
9.689 | -103.7393
9.6925 | -103.7527
9.6961 | -103.7704
9.6997 | -103.7855
9.7032 | -103.7991
9.7068 | -103.8192
9.7104 | -103.8416
9.714 | -103.8673
9.7175 | -103.8926
9.7211 | -103.916
9.7247 | -103.9416
9.7282 | -103.9659
9.7318 | -103.9879
9.7354 | -104.0129
9.7389 | -104.0388
9.7425 | -104.0633
9.7461 | -104.084
9.7497 | -104.103
9.7532 | -104.1187
9.7568 | -104.133
9.7604 | -104.1459
9.7639 | -104.1556
9.7675 | -104.1585
9.7711 | -104.1603
9.7746 | -104.1621
9.7782 | -104.1604
9.7818 | -104.1562
9.7854 | -104.1536
9.7889 | -104.1518
9.7925 | -104.1458
9.7961 | -104.1386
9.7996 | -104.13
9.8032 | -104.1249
9.8068 | -104.1218
9.8103 | -104.1175
9.8139 | -104.115
9.8175 | -104.1112
9.8211 | -104.1079
9.8246 | -104.1036
9.8282 | -104.0985
9.8318 | -104.095
9.8353 | -104.0886
9.8389 | -104.0797
9.8425 | -104.0733
9.846 | -104.0668
9.8496 | -104.0587
9.8532 | -104.0503
9.8568 | -104.0423
9.8603 | -104.0423
9.8639 | -104.0439
9.8675 | -104.0423
9.871 | -104.0418
9.8746 | -104.0361
9.8782 | -104.0334
9.8817 | -104.0295
9.8853 | -104.0262
9.8889 | -104.0239
9.8925 | -104.0211
9.896 | -104.0167
9.8996 | -104.0058
9.9032 | -103.9938
9.9067 | -103.9839
9.9103 | -103.975
9.9139 | -103.9629
9.9174 | -103.948
9.921 | -103.9303
9.9246 | -103.9125
9.9282 | -103.8903
9.9317 | -103.8722
9.9353 | -103.8587
9.9389 | -103.8468
9.9424 | -103.8404
9.946 | -103.8355
9.9496 | -103.8306
9.9531 | -103.8317
9.9567 | -103.8334
9.9603 | -103.8394
9.9639 | -103.8459
9.9674 | -103.8525
9.971 | -103.8636
9.9746 | -103.8764
9.9781 | -103.89
9.9817 | -103.9028
9.9853 | -103.9131
9.9888 | -103.9247
9.9924 | -103.9349
9.996 | -103.9408
9.9996 | -103.9413
10.0031 | -103.9437
10.0067 | -103.9435
10.0103 | -103.9413
10.0138 | -103.9423
10.0174 | -103.9424
10.021 | -103.948
10.0245 | -103.9492
10.0281 | -103.95
10.0317 | -103.9542
10.0353 | -103.9587
10.0388 | -103.9642
10.0424 | -103.9718
10.046 | -103.9854
10.0495 | -103.9983
10.0531 | -104.011
10.0567 | -104.0255
10.0602 | -104.0467
10.0638 | -104.0665
10.0674 | -104.0862
10.071 | -104.1075
10.0745 | -104.1286
10.0781 | -104.1442
10.0817 | -104.1539
10.0852 | -104.1664
10.0888 | -104.1716
10.0924 | -104.1703
10.0959 | -104.1623
10.0995 | -104.1507
10.1031 | -104.1359
10.1067 | -104.116
10.1102 | -104.0955
10.1138 | -104.0726
10.1174 | -104.0438
10.1209 | -104.012
10.1245 | -103.983
10.1281 | -103.954
10.1316 | -103.9254
10.1352 | -103.8947
10.1388 | -103.8697
10.1424 | -103.8492
10.1459 | -103.838
10.1495 | -103.829
10.1531 | -103.8219
10.1566 | -103.8248
10.1602 | -103.8341
10.1638 | -103.8459
10.1673 | -103.8575
10.1709 | -103.8723
10.1745 | -103.8895
10.1781 | -103.9084
10.1816 | -103.9319
10.1852 | -103.9531
10.1888 | -103.9741
10.1923 | -103.9889
10.1959 | -104.0009
10.1995 | -104.0078
10.203 | -104.0109
10.2066 | -104.0061
10.2102 | -104.0019
10.2138 | -103.9956
10.2173 | -103.9826
10.2209 | -103.9666
10.2245 | -103.9495
10.228 | -103.9304
10.2316 | -103.9093
10.2352 | -103.8892
10.2387 | -103.8729
10.2423 | -103.8582
10.2459 | -103.8439
10.2495 | -103.837
10.253 | -103.8341
10.2566 | -103.8381
10.2602 | -103.847
10.2637 | -103.8618
10.2673 | -103.8828
10.2709 | -103.9076
10.2744 | -103.9407
10.278 | -103.9774
10.2816 | -104.0215
10.2852 | -104.0663
10.2887 | -104.1108
10.2923 | -104.1547
10.2959 | -104.1978
10.2994 | -104.2243
10.303 | -104.2599
10.3066 | -104.2937
10.3101 | -104.3256
10.3137 | -104.3573
10.3173 | -104.387
10.3209 | -104.4158
10.3244 | -104.4405
10.328 | -104.463
10.3316 | -104.4806
10.3351 | -104.4975
10.3387 | -104.5117
10.3423 | -104.5263
10.3458 | -104.5384
10.3494 | -104.5509
10.353 | -104.5636
10.3566 | -104.5759
10.3601 | -104.5876
10.3637 | -104.595
10.3673 | -104.6028
10.3708 | -104.6097
10.3744 | -104.6117
10.378 | -104.6157
10.3815 | -104.6161
10.3851 | -104.6143
10.3887 | -104.6132
10.3923 | -104.6087
10.3958 | -104.596
10.3994 | -104.5758
10.403 | -104.5522
10.4065 | -104.5242
10.4101 | -104.4957
10.4137 | -104.4619
10.4172 | -104.4254
10.4208 | -104.3902
10.4244 | -104.3542
10.428 | -104.3183
10.4315 | -104.2798
10.4351 | -104.2408
10.4387 | -104.2067
10.4422 | -104.1755
10.4458 | -104.148
10.4494 | -104.123
10.4529 | -104.0992
10.4565 | -104.0758
10.4601 | -104.056
10.4637 | -104.0399
10.4672 | -104.0277
10.4708 | -104.0175
10.4744 | -104.0079
10.4779 | -104.0035
10.4815 | -104.0008
10.4851 | -104.0016
10.4886 | -104.0064
10.4922 | -104.0124
10.4958 | -104.0174
10.4994 | -104.0247
10.5029 | -104.0344
10.5065 | -104.0463
10.5101 | -104.0584
10.5136 | -104.0716
10.5172 | -104.092
10.5208 | -104.1123
10.5243 | -104.1351
10.5279 | -104.1626
10.5315 | -104.1871
10.5351 | -104.2129
10.5386 | -104.2395
10.5422 | -104.2635
10.5458 | -104.2862
10.5493 | -104.3032
10.5529 | -104.3208
10.5565 | -104.3383
10.56 | -104.3532
10.5636 | -104.3637
10.5672 | -104.3713
10.5708 | -104.3753
10.5743 | -104.3758
10.5779 | -104.372
10.5815 | -104.3633
10.585 | -104.3531
10.5886 | -104.3415
10.5922 | -104.3243
10.5957 | -104.3042
10.5993 | -104.2835
10.6029 | -104.2675
10.6065 | -104.2555
10.61 | -104.2421
10.6136 | -104.2238
10.6172 | -104.2052
10.6207 | -104.1859
10.6243 | -104.1739
10.6279 | -104.166
10.6314 | -104.1596
10.635 | -104.157
10.6386 | -104.1542
10.6422 | -104.1567
10.6457 | -104.1634
10.6493 | -104.1697
10.6529 | -104.1787
10.6564 | -104.1901
10.66 | -104.2
10.6636 | -104.2133
10.6671 | -104.2301
10.6707 | -104.2455
10.6743 | -104.2655
10.6779 | -104.2865
10.6814 | -104.3057
10.685 | -104.3248
10.6886 | -104.3454
10.6921 | -104.3661
10.6957 | -104.3868
10.6993 | -104.4085
10.7028 | -104.4318
10.7064 | -104.4508
10.71 | -104.4655
10.7136 | -104.48
10.7171 | -104.4924
10.7207 | -104.5029
10.7243 | -104.5117
10.7278 | -104.5117
10.7314 | -104.5146
10.735 | -104.5194
10.7385 | -104.5235
10.7421 | -104.5259
10.7457 | -104.5273
10.7493 | -104.5235
10.7528 | -104.5139
10.7564 | -104.4998
10.76 | -104.4829
10.7635 | -104.4686
10.7671 | -104.4576
10.7707 | -104.4467
10.7742 | -104.438
10.7778 | -104.425
10.7814 | -104.4098
10.785 | -104.3918
10.7885 | -104.374
10.7921 | -104.3608
10.7957 | -104.3498
10.7992 | -104.3376
10.8028 | -104.3297
10.8064 | -104.3201
10.8099 | -104.3084
10.8135 | -104.3041
10.8171 | -104.2985
10.8207 | -104.2897
10.8242 | -104.2757
10.8278 | -104.2652
10.8314 | -104.2588
10.8349 | -104.2523
10.8385 | -104.2485
10.8421 | -104.2484
10.8456 | -104.2435
10.8492 | -104.2403
10.8528 | -104.2422
10.8564 | -104.242
10.8599 | -104.244
10.8635 | -104.2472
10.8671 | -104.2508
10.8706 | -104.2562
10.8742 | -104.2659
10.8778 | -104.2785
10.8813 | -104.2938
10.8849 | -104.3124
10.8885 | -104.329
10.8921 | -104.3478
10.8956 | -104.3666
10.8992 | -104.3847
10.9028 | -104.4048
10.9063 | -104.429
10.9099 | -104.458
10.9135 | -104.4861
10.917 | -104.5113
10.9206 | -104.5299
10.9242 | -104.5498
10.9278 | -104.5683
10.9313 | -104.5853
10.9349 | -104.5979
10.9385 | -104.6134
10.942 | -104.6269
10.9456 | -104.6356
10.9492 | -104.6386
10.9527 | -104.6296
10.9563 | -104.6229
10.9599 | -104.6165
10.9635 | -104.6082
10.967 | -104.5973
10.9706 | -104.5824
10.9742 | -104.5658
10.9777 | -104.5501
10.9813 | -104.5368
10.9849 | -104.525
10.9884 | -104.5122
10.992 | -104.4941
10.9956 | -104.4796
10.9992 | -104.4637
11.0027 | -104.444
11.0063 | -104.4301
11.0099 | -104.4178
11.0134 | -104.4081
11.017 | -104.4028
11.0206 | -104.4009
11.0241 | -104.399
11.0277 | -104.3972
11.0313 | -104.3965
11.0349 | -104.3988
11.0384 | -104.4035
11.042 | -104.4082
11.0456 | -104.4139
11.0491 | -104.4212
11.0527 | -104.4284
11.0563 | -104.4394
11.0598 | -104.4532
11.0634 | -104.4659
11.067 | -104.4771
11.0706 | -104.4842
11.0741 | -104.4947
11.0777 | -104.5085
11.0813 | -104.524
11.0848 | -104.5345
11.0884 | -104.5458
11.092 | -104.56
11.0955 | -104.5737
11.0991 | -104.5836
11.1027 | -104.5959
11.1063 | -104.6088
11.1098 | -104.6235
11.1134 | -104.6406
11.117 | -104.6548
11.1205 | -104.6625
11.1241 | -104.6685
11.1277 | -104.679
11.1312 | -104.6897
11.1348 | -104.6978
11.1384 | -104.7045
11.142 | -104.711
11.1455 | -104.716
11.1491 | -104.719
11.1527 | -104.7209
11.1562 | -104.7196
11.1598 | -104.7218
11.1634 | -104.7274
11.1669 | -104.7293
11.1705 | -104.7278
11.1741 | -104.7197
11.1777 | -104.716
11.1812 | -104.712
11.1848 | -104.7073
11.1884 | -104.706
11.1919 | -104.7053
11.1955 | -104.7015
11.1991 | -104.6993
11.2026 | -104.6982
11.2062 | -104.7009
11.2098 | -104.7017
11.2134 | -104.7006
11.2169 | -104.6951
11.2205 | -104.6871
11.2241 | -104.6797
11.2276 | -104.6689
11.2312 | -104.6573
11.2348 | -104.6409
11.2383 | -104.6203
11.2419 | -104.6003
11.2455 | -104.5816
11.2491 | -104.5598
11.2526 | -104.5348
11.2562 | -104.5144
11.2598 | -104.498
11.2633 | -104.4818
11.2669 | -104.4672
11.2705 | -104.4546
11.274 | -104.4412
11.2776 | -104.4259
11.2812 | -104.4156
11.2848 | -104.4085
11.2883 | -104.3993
11.2919 | -104.3911
11.2955 | -104.389
11.299 | -104.3883
11.3026 | -104.3893
11.3062 | -104.3919
11.3097 | -104.3968
11.3133 | -104.4072
11.3169 | -104.4194
11.3205 | -104.4267
11.324 | -104.4334
11.3276 | -104.4421
11.3312 | -104.4502
11.3347 | -104.4592
11.3383 | -104.4708
11.3419 | -104.479
11.3454 | -104.4881
11.349 | -104.4984
11.3526 | -104.509
11.3562 | -104.5183
11.3597 | -104.5243
11.3633 | -104.5251
11.3669 | -104.5278
11.3704 | -104.5332
11.374 | -104.5375
11.3776 | -104.5407
11.3811 | -104.5437
11.3847 | -104.5526
11.3883 | -104.5579
11.3919 | -104.563
11.3954 | -104.5667
11.399 | -104.5714
11.4026 | -104.5745
11.4061 | -104.5748
11.4097 | -104.5763
11.4133 | -104.5787
11.4168 | -104.5817
11.4204 | -104.584
11.424 | -104.5822
11.4276 | -104.5796
11.4311 | -104.5788
11.4347 | -104.5756
11.4383 | -104.573
11.4418 | -104.577
11.4454 | -104.5807
11.449 | -104.5793
11.4525 | -104.579
11.4561 | -104.5792
11.4597 | -104.5708
11.4633 | -104.5641
11.4668 | -104.5594
11.4704 | -104.5554
11.474 | -104.5521
11.4775 | -104.5443
11.4811 | -104.5371
11.4847 | -104.5358
11.4882 | -104.5313
11.4918 | -104.5274
11.4954 | -104.5245
11.499 | -104.5206
11.5025 | -104.5162
11.5061 | -104.5084
11.5097 | -104.4977
11.5132 | -104.4887
11.5168 | -104.4802
11.5204 | -104.4701
11.5239 | -104.4632
11.5275 | -104.4595
11.5311 | -104.4599
11.5347 | -104.4564
11.5382 | -104.4504
11.5418 | -104.4455
11.5454 | -104.4469
11.5489 | -104.4505
11.5525 | -104.4479
11.5561 | -104.448
11.5596 | -104.4508
11.5632 | -104.4565
11.5668 | -104.462
11.5704 | -104.4695
11.5739 | -104.4765
11.5775 | -104.4813
11.5811 | -104.4873
11.5846 | -104.4923
11.5882 | -104.4933
11.5918 | -104.4963
11.5953 | -104.5002
11.5989 | -104.5033
11.6025 | -104.5052
11.6061 | -104.5065
11.6096 | -104.5065
11.6132 | -104.5066
11.6168 | -104.5118
11.6203 | -104.5158
11.6239 | -104.515
11.6275 | -104.5196
11.631 | -104.525
11.6346 | -104.5307
11.6382 | -104.5383
11.6418 | -104.5468
11.6453 | -104.5564
11.6489 | -104.5691
11.6525 | -104.5837
11.656 | -104.6004
11.6596 | -104.6099
11.6632 | -104.6171
11.6667 | -104.6275
11.6703 | -104.6391
11.6739 | -104.6532
11.6775 | -104.668
11.681 | -104.6788
11.6846 | -104.6877
11.6882 | -104.6951
11.6917 | -104.6987
11.6953 | -104.6998
11.6989 | -104.6995
11.7024 | -104.7013
11.706 | -104.7016
11.7096 | -104.7018
11.7132 | -104.6993
11.7167 | -104.6954
11.7203 | -104.6944
11.7239 | -104.6965
11.7274 | -104.6966
11.731 | -104.692
11.7346 | -104.6867
11.7381 | -104.6833
11.7417 | -104.6833
11.7453 | -104.6831
11.7489 | -104.6826
11.7524 | -104.6845
11.756 | -104.6855
11.7596 | -104.6863
11.7631 | -104.6893
11.7667 | -104.6911
11.7703 | -104.6953
11.7738 | -104.7007
11.7774 | -104.7005
11.781 | -104.7019
11.7846 | -104.703
11.7881 | -104.7065
11.7917 | -104.7141
11.7953 | -104.7206
11.7988 | -104.7199
11.8024 | -104.7155
11.806 | -104.7133
11.8095 | -104.7116
11.8131 | -104.7139
11.8167 | -104.7169
11.8203 | -104.7195
11.8238 | -104.724
11.8274 | -104.7329
11.831 | -104.7444
11.8345 | -104.7531
11.8381 | -104.7627
11.8417 | -104.7703
11.8452 | -104.7763
11.8488 | -104.785
11.8524 | -104.794
11.856 | -104.8026
11.8595 | -104.8109
11.8631 | -104.8169
11.8667 | -104.823
11.8702 | -104.8277
11.8738 | -104.8311
11.8774 | -104.8366
11.8809 | -104.8429
11.8845 | -104.8457
11.8881 | -104.85
11.8917 | -104.8551
11.8952 | -104.8604
11.8988 | -104.8652
11.9024 | -104.8712
11.9059 | -104.8768
11.9095 | -104.882
11.9131 | -104.8882
11.9166 | -104.8921
11.9202 | -104.8931
11.9238 | -104.8928
11.9274 | -104.8974
11.9309 | -104.9014
11.9345 | -104.9013
11.9381 | -104.8992
11.9416 | -104.8985
11.9452 | -104.8973
11.9488 | -104.8966
11.9523 | -104.8956
11.9559 | -104.8903
11.9595 | -104.8808
11.9631 | -104.8742
11.9666 | -104.8695
11.9702 | -104.8674
11.9738 | -104.8635
11.9773 | -104.858
11.9809 | -104.8523
11.9845 | -104.8471
11.988 | -104.8416
11.9916 | -104.8328
11.9952 | -104.8253
11.9988 | -104.8169
12.0023 | -104.8064
12.0059 | -104.7986
12.0095 | -104.7871
12.013 | -104.7781
12.0166 | -104.773
12.0202 | -104.7663
12.0237 | -104.7568
12.0273 | -104.7475
12.0309 | -104.739
12.0344 | -104.7296
12.038 | -104.7235
12.0416 | -104.7163
12.0452 | -104.7109
12.0487 | -104.7068
12.0523 | -104.7004
12.0559 | -104.6979
12.0594 | -104.7006
12.063 | -104.7061
12.0666 | -104.7117
12.0701 | -104.7158
12.0737 | -104.7245
12.0773 | -104.7369
12.0809 | -104.7503
12.0844 | -104.7678
12.088 | -104.7831
12.0916 | -104.8018
12.0951 | -104.8248
12.0987 | -104.848
12.1023 | -104.8719
12.1059 | -104.8957
12.1094 | -104.9187
12.113 | -104.9447
12.1166 | -104.964
12.1201 | -104.9847
12.1237 | -105.002
12.1273 | -105.0163
12.1308 | -105.0316
12.1344 | -105.0447
12.138 | -105.05
12.1416 | -105.0464
12.1451 | -105.0404
12.1487 | -105.0308
12.1523 | -105.0186
12.1558 | -105.0013
12.1594 | -104.9765
12.163 | -104.9473
12.1665 | -104.9192
12.1701 | -104.8925
12.1737 | -104.8644
12.1773 | -104.8313
12.1808 | -104.796
12.1844 | -104.762
12.188 | -104.7268
12.1915 | -104.6937
12.1951 | -104.6608
12.1987 | -104.6332
12.2022 | -104.6079
12.2058 | -104.5858
12.2094 | -104.5684
12.2129 | -104.5515
12.2165 | -104.5386
12.2201 | -104.5265
12.2237 | -104.5186
12.2272 | -104.5142
12.2308 | -104.512
12.2344 | -104.5145
12.2379 | -104.5191
12.2415 | -104.5263
12.2451 | -104.5396
12.2486 | -104.5548
12.2522 | -104.5711
12.2558 | -104.5878
12.2594 | -104.6043
12.2629 | -104.625
12.2665 | -104.6485
12.2701 | -104.6715
12.2736 | -104.6969
12.2772 | -104.7207
12.2808 | -104.7435
12.2843 | -104.769
12.2879 | -104.7932
12.2915 | -104.8205
12.2951 | -104.8402
12.2986 | -104.8605
12.3022 | -104.877
12.3058 | -104.8952
12.3093 | -104.9105
12.3129 | -104.9216
12.3165 | -104.9351
12.3201 | -104.948
12.3236 | -104.9605
12.3272 | -104.9669
12.3308 | -104.9711
12.3343 | -104.9741
12.3379 | -104.9758
12.3415 | -104.9741
12.345 | -104.9679
12.3486 | -104.9607
12.3522 | -104.9524
12.3558 | -104.9415
12.3593 | -104.9198
12.3629 | -104.9065
12.3665 | -104.8916
12.37 | -104.872
12.3736 | -104.8538
12.3772 | -104.8354
12.3807 | -104.8186
12.3843 | -104.7999
12.3879 | -104.7829
12.3915 | -104.7662
12.395 | -104.7484
12.3986 | -104.7344
12.4022 | -104.7194
12.4057 | -104.7037
12.4093 | -104.6889
12.4129 | -104.6734
12.4164 | -104.6624
12.42 | -104.6518
12.4236 | -104.6407
12.4271 | -104.6323
12.4307 | -104.625
12.4343 | -104.6186
12.4379 | -104.6101
12.4414 | -104.6029
12.445 | -104.5958
12.4486 | -104.5889
12.4521 | -104.5824
12.4557 | -104.5787
12.4593 | -104.5776
12.4628 | -104.5796
12.4664 | -104.5821
12.47 | -104.5861
12.4736 | -104.5937
12.4771 | -104.6017
12.4807 | -104.6072
12.4843 | -104.6103
12.4878 | -104.6145
12.4914 | -104.6199
12.495 | -104.6237
12.4985 | -104.626
12.5021 | -104.6291
12.5057 | -104.6336
12.5093 | -104.6385
12.5128 | -104.6456
12.5164 | -104.6506
12.52 | -104.6565
12.5235 | -104.6595
12.5271 | -104.6623
12.5307 | -104.6681
12.5343 | -104.6717
12.5378 | -104.6741
12.5414 | -104.6781
12.545 | -104.6794
12.5485 | -104.6807
12.5521 | -104.6838
12.5557 | -104.6863
12.5592 | -104.6887
12.5628 | -104.6885
12.5664 | -104.6903
12.57 | -104.6914
12.5735 | -104.6916
12.5771 | -104.6912
12.5807 | -104.6894
12.5842 | -104.6901
12.5878 | -104.6942
12.5914 | -104.6931
12.5949 | -104.689
12.5985 | -104.6838
12.6021 | -104.6807
12.6057 | -104.6783
12.6092 | -104.6749
12.6128 | -104.6723
12.6164 | -104.6664
12.6199 | -104.663
12.6235 | -104.6593
12.6271 | -104.6545
12.6306 | -104.6504
12.6342 | -104.6455
12.6378 | -104.6426
12.6413 | -104.6413
12.6449 | -104.6365
12.6485 | -104.6338
12.6521 | -104.6347
12.6556 | -104.6392
12.6592 | -104.6451
12.6628 | -104.6554
12.6663 | -104.6685
12.6699 | -104.6839
12.6735 | -104.702
12.677 | -104.7186
12.6806 | -104.7366
12.6842 | -104.757
12.6878 | -104.7816
12.6913 | -104.809
12.6949 | -104.8353
12.6985 | -104.8642
12.702 | -104.8917
12.7056 | -104.9222
12.7092 | -104.9534
12.7127 | -104.9816
12.7163 | -105.0098
12.7199 | -105.0356
12.7235 | -105.059
12.727 | -105.0841
12.7306 | -105.1081
12.7342 | -105.1254
12.7377 | -105.1421
12.7413 | -105.1575
12.7449 | -105.1719
12.7484 | -105.1816
12.752 | -105.1877
12.7556 | -105.1964
12.7592 | -105.2052
12.7627 | -105.2072
12.7663 | -105.2013
12.7699 | -105.1976
12.7734 | -105.1988
12.777 | -105.1994
12.7806 | -105.197
12.7841 | -105.1936
12.7877 | -105.1877
12.7913 | -105.1794
12.7949 | -105.173
12.7984 | -105.1717
12.802 | -105.168
12.8056 | -105.1654
12.8091 | -105.1657
12.8127 | -105.1655
12.8163 | -105.1664
12.8198 | -105.1657
12.8234 | -105.1663
12.827 | -105.1718
12.8306 | -105.1762
12.8341 | -105.1784
12.8377 | -105.1832
12.8413 | -105.1882
12.8448 | -105.1957
12.8484 | -105.2012
12.852 | -105.207
12.8555 | -105.2112
12.8591 | -105.2192
12.8627 | -105.2286
12.8663 | -105.2327
12.8698 | -105.2333
12.8734 | -105.2272
12.877 | -105.2245
12.8805 | -105.223
12.8841 | -105.2208
12.8877 | -105.2168
12.8912 | -105.2093
12.8948 | -105.2026
12.8984 | -105.1959
12.902 | -105.1897
12.9055 | -105.182
12.9091 | -105.1744
12.9127 | -105.1641
12.9162 | -105.1541
12.9198 | -105.1472
12.9234 | -105.1368
12.9269 | -105.1304
12.9305 | -105.1283
12.9341 | -105.1299
12.9377 | -105.1311
12.9412 | -105.1312
12.9448 | -105.1294
12.9484 | -105.1306
12.9519 | -105.136
12.9555 | -105.1377
12.9591 | -105.1389
12.9626 | -105.1383
12.9662 | -105.1384
12.9698 | -105.1386
12.9734 | -105.14
12.9769 | -105.1399
12.9805 | -105.1341
12.9841 | -105.1253
12.9876 | -105.116
12.9912 | -105.1024
12.9948 | -105.0898
12.9983 | -105.0794
13.0019 | -105.07
13.0055 | -105.0584
13.0091 | -105.0454
13.0126 | -105.0316
13.0162 | -105.0222
13.0198 | -105.014
13.0233 | -105.0025
13.0269 | -104.9942
13.0305 | -104.9869
13.034 | -104.9799
13.0376 | -104.9727
13.0412 | -104.9702
13.0448 | -104.9673
13.0483 | -104.9627
13.0519 | -104.9659
13.0555 | -104.9714
13.059 | -104.9747
13.0626 | -104.979
13.0662 | -104.9812
13.0697 | -104.9857
13.0733 | -104.9896
13.0769 | -104.9979
13.0805 | -105.0075
13.084 | -105.0143
13.0876 | -105.0205
13.0912 | -105.0214
13.0947 | -105.0252
13.0983 | -105.0268
13.1019 | -105.0256
13.1054 | -105.024
13.109 | -105.0239
13.1126 | -105.0231
13.1162 | -105.0166
13.1197 | -105.007
13.1233 | -104.9952
13.1269 | -104.9836
13.1304 | -104.971
13.134 | -104.9591
13.1376 | -104.9475
13.1411 | -104.9364
13.1447 | -104.9218
13.1483 | -104.9058
13.1519 | -104.8937
13.1554 | -104.8811
13.159 | -104.8656
13.1626 | -104.8521
13.1661 | -104.8418
13.1697 | -104.8352
13.1733 | -104.8287
13.1768 | -104.8222
13.1804 | -104.8175
13.184 | -104.8107
13.1876 | -104.8074
13.1911 | -104.805
13.1947 | -104.8057
13.1983 | -104.8089
13.2018 | -104.8132
13.2054 | -104.8165
13.209 | -104.8216
13.2125 | -104.8286
13.2161 | -104.8321
13.2197 | -104.8355
13.2233 | -104.8404
13.2268 | -104.8458
13.2304 | -104.848
13.234 | -104.8522
13.2375 | -104.8575
13.2411 | -104.8588
13.2447 | -104.8581
13.2482 | -104.8639
13.2518 | -104.8694
13.2554 | -104.8743
13.259 | -104.8788
13.2625 | -104.8837
13.2661 | -104.8869
13.2697 | -104.8908
13.2732 | -104.8974
13.2768 | -104.9066
13.2804 | -104.919
13.2839 | -104.9322
13.2875 | -104.9458
13.2911 | -104.9585
13.2947 | -104.9698
13.2982 | -104.9799
13.3018 | -104.9892
13.3054 | -104.998
13.3089 | -105.0082
13.3125 | -105.0173
13.3161 | -105.025
13.3196 | -105.032
13.3232 | -105.0366
13.3268 | -105.0374
13.3304 | -105.0327
13.3339 | -105.0252
13.3375 | -105.0173
13.3411 | -105.007
13.3446 | -104.9903
13.3482 | -104.9708
13.3518 | -104.9486
13.3553 | -104.9255
13.3589 | -104.8989
13.3625 | -104.8673
13.3661 | -104.8343
13.3696 | -104.7998
13.3732 | -104.7628
13.3768 | -104.7207
13.3803 | -104.6783
13.3839 | -104.6355
13.3875 | -104.5973
13.391 | -104.5605
13.3946 | -104.526
13.3982 | -104.4934
13.4018 | -104.4604
13.4053 | -104.433
13.4089 | -104.4122
13.4125 | -104.3935
13.416 | -104.3773
13.4196 | -104.3651
13.4232 | -104.3579
13.4267 | -104.3542
13.4303 | -104.355
13.4339 | -104.3605
13.4375 | -104.3711
13.441 | -104.3839
13.4446 | -104.4012
13.4482 | -104.4233
13.4517 | -104.4467
13.4553 | -104.4728
13.4589 | -104.5027
13.4624 | -104.5353
13.466 | -104.571
13.4696 | -104.6065
13.4732 | -104.6452
13.4767 | -104.6837
13.4803 | -104.7227
13.4839 | -104.7642
13.4874 | -104.8048
13.491 | -104.8402
13.4946 | -104.876
13.4981 | -104.9117
13.5017 | -104.9422
13.5053 | -104.9684
13.5089 | -104.9911
13.5124 | -105.0109
13.516 | -105.0268
13.5196 | -105.0411
13.5231 | -105.0559
13.5267 | -105.0689
13.5303 | -105.0747
13.5338 | -105.0752
13.5374 | -105.0759
13.541 | -105.0778
13.5446 | -105.0738
13.5481 | -105.0656
13.5517 | -105.0586
13.5553 | -105.0538
13.5588 | -105.049
13.5624 | -105.0452
13.566 | -105.0424
13.5695 | -105.0379
13.5731 | -105.0328
13.5767 | -105.0307
13.5803 | -105.028
13.5838 | -105.0244
13.5874 | -105.0232
13.591 | -105.0229
13.5945 | -105.0281
13.5981 | -105.0365
13.6017 | -105.0438
13.6052 | -105.0511
13.6088 | -105.0604
13.6124 | -105.0733
13.616 | -105.0887
13.6195 | -105.1048
13.6231 | -105.1223
13.6267 | -105.1411
13.6302 | -105.1639
13.6338 | -105.1858
13.6374 | -105.2066
13.6409 | -105.2273
13.6445 | -105.2489
13.6481 | -105.2705
13.6517 | -105.2926
13.6552 | -105.3127
13.6588 | -105.3316
13.6624 | -105.3451
13.6659 | -105.3579
13.6695 | -105.3661
13.6731 | -105.3695
13.6766 | -105.3731
13.6802 | -105.374
13.6838 | -105.3727
13.6874 | -105.3694
13.6909 | -105.3589
13.6945 | -105.3423
13.6981 | -105.3294
13.7016 | -105.3157
13.7052 | -105.2966
13.7088 | -105.274
13.7123 | -105.2515
13.7159 | -105.2272
13.7195 | -105.2008
13.7231 | -105.1748
13.7266 | -105.1505
13.7302 | -105.1269
13.7338 | -105.1039
13.7373 | -105.0795
13.7409 | -105.0577
13.7445 | -105.0402
13.748 | -105.0233
13.7516 | -105.0075
13.7552 | -104.9941
13.7588 | -104.9832
13.7623 | -104.9772
13.7659 | -104.9776
13.7695 | -104.981
13.773 | -104.9854
13.7766 | -104.994
13.7802 | -105.004
13.7837 | -105.0177
13.7873 | -105.0358
13.7909 | -105.0548
13.7945 | -105.0754
13.798 | -105.0985
13.8016 | -105.1255
13.8052 | -105.155
13.8087 | -105.1843
13.8123 | -105.2138
13.8159 | -105.2414
13.8194 | -105.2687
13.823 | -105.2942
13.8266 | -105.3176
13.8302 | -105.3422
13.8337 | -105.3656
13.8373 | -105.3896
13.8409 | -105.4117
13.8444 | -105.4311
13.848 | -105.4469
13.8516 | -105.4595
13.8551 | -105.4687
13.8587 | -105.4796
13.8623 | -105.4859
13.8659 | -105.4934
13.8694 | -105.5027
13.873 | -105.5083
13.8766 | -105.5141
13.8801 | -105.5202
13.8837 | -105.5281
13.8873 | -105.5383
13.8908 | -105.5472
13.8944 | -105.5516
13.898 | -105.5571
13.9016 | -105.5666
13.9051 | -105.5787
13.9087 | -105.5875
13.9123 | -105.5956
13.9158 | -105.606
13.9194 | -105.6153
13.923 | -105.6223
13.9265 | -105.6303
13.9301 | -105.6385
13.9337 | -105.647
13.9373 | -105.654
13.9408 | -105.6607
13.9444 | -105.6654
13.948 | -105.6709
13.9515 | -105.6775
13.9551 | -105.6804
13.9587 | -105.6784
13.9622 | -105.6734
13.9658 | -105.6691
13.9694 | -105.6649
13.973 | -105.6599
13.9765 | -105.652
13.9801 | -105.6422
13.9837 | -105.6277
13.9872 | -105.6122
13.9908 | -105.5953
13.9944 | -105.5792
13.9979 | -105.5611
14.0015 | -105.5366
14.0051 | -105.5087
14.0087 | -105.4825
14.0122 | -105.4542
14.0158 | -105.427
14.0194 | -105.3995
14.0229 | -105.3728
14.0265 | -105.3436
14.0301 | -105.3139
14.0336 | -105.2854
14.0372 | -105.2595
14.0408 | -105.234
14.0444 | -105.2098
14.0479 | -105.1875
14.0515 | -105.1676
14.0551 | -105.1537
14.0586 | -105.1388
14.0622 | -105.1279
14.0658 | -105.1164
14.0693 | -105.1064
14.0729 | -105.1018
14.0765 | -105.0963
14.0801 | -105.094
14.0836 | -105.0942
14.0872 | -105.0967
14.0908 | -105.101
14.0943 | -105.1047
14.0979 | -105.1089
14.1015 | -105.1146
14.105 | -105.1222
14.1086 | -105.1288
14.1122 | -105.1364
14.1158 | -105.1443
14.1193 | -105.1496
14.1229 | -105.1563
14.1265 | -105.1643
14.13 | -105.1702
14.1336 | -105.1748
14.1372 | -105.178
14.1407 | -105.1844
14.1443 | -105.1899
14.1479 | -105.1929
14.1515 | -105.1967
14.155 | -105.1983
14.1586 | -105.2012
14.1622 | -105.2089
14.1657 | -105.2155
14.1693 | -105.2194
14.1729 | -105.2258
14.1764 | -105.2333
14.18 | -105.2417
14.1836 | -105.2533
14.1872 | -105.267
14.1907 | -105.2807
14.1943 | -105.2954
14.1979 | -105.3105
14.2014 | -105.3261
14.205 | -105.3463
14.2086 | -105.3659
14.2121 | -105.3816
14.2157 | -105.3996
14.2193 | -105.4171
14.2229 | -105.4342
14.2264 | -105.4521
14.23 | -105.4704
14.2336 | -105.4862
14.2371 | -105.502
14.2407 | -105.5166
14.2443 | -105.531
14.2478 | -105.5464
14.2514 | -105.5651
14.255 | -105.5835
14.2586 | -105.5988
14.2621 | -105.6161
14.2657 | -105.6313
14.2693 | -105.646
14.2728 | -105.6607
14.2764 | -105.6712
14.28 | -105.6809
14.2835 | -105.6869
14.2871 | -105.6919
14.2907 | -105.695
14.2943 | -105.6975
14.2978 | -105.7006
14.3014 | -105.7013
14.305 | -105.6989
14.3085 | -105.6918
14.3121 | -105.6851
14.3157 | -105.6776
14.3192 | -105.6698
14.3228 | -105.6608
14.3264 | -105.6519
14.33 | -105.6422
14.3335 | -105.6338
14.3371 | -105.6265
14.3407 | -105.6188
14.3442 | -105.6091
14.3478 | -105.6037
14.3514 | -105.6011
14.3549 | -105.596
14.3585 | -105.5967
14.3621 | -105.6015
14.3657 | -105.6048
14.3692 | -105.6083
14.3728 | -105.6119
14.3764 | -105.6145
14.3799 | -105.6193
14.3835 | -105.6265
14.3871 | -105.6333
14.3906 | -105.6414
14.3942 | -105.6521
14.3978 | -105.6631
14.4014 | -105.6762
14.4049 | -105.6908
14.4085 | -105.7031
14.4121 | -105.7173
14.4156 | -105.732
14.4192 | -105.7361
14.4228 | -105.7444
14.4263 | -105.7535
14.4299 | -105.7602
14.4335 | -105.7639
14.4371 | -105.7654
14.4406 | -105.7693
14.4442 | -105.7692
14.4478 | -105.767
14.4513 | -105.7646
14.4549 | -105.7591
14.4585 | -105.751
14.462 | -105.7438
14.4656 | -105.7353
14.4692 | -105.724
14.4728 | -105.714
14.4763 | -105.7065
14.4799 | -105.6948
14.4835 | -105.6828
14.487 | -105.6708
14.4906 | -105.6606
14.4942 | -105.6522
14.4977 | -105.6415
14.5013 | -105.6299
14.5049 | -105.6183
14.5085 | -105.6069
14.512 | -105.5955
14.5156 | -105.5843
14.5192 | -105.5717
14.5227 | -105.5586
14.5263 | -105.5457
14.5299 | -105.5331
14.5334 | -105.5205
14.537 | -105.5123
14.5406 | -105.5039
14.5442 | -105.4955
14.5477 | -105.4882
14.5513 | -105.4798
14.5549 | -105.4696
14.5584 | -105.4613
14.562 | -105.4554
14.5656 | -105.4487
14.5691 | -105.4405
14.5727 | -105.4361
14.5763 | -105.4341
14.5799 | -105.4318
14.5834 | -105.4291
14.587 | -105.4265
14.5906 | -105.4254
14.5941 | -105.4258
14.5977 | -105.4238
14.6013 | -105.4211
14.6048 | -105.42
14.6084 | -105.4202
14.612 | -105.4244
14.6156 | -105.4282
14.6191 | -105.4325
14.6227 | -105.4396
14.6263 | -105.4498
14.6298 | -105.4604
14.6334 | -105.4681
14.637 | -105.4691
14.6405 | -105.4733
14.6441 | -105.4813
14.6477 | -105.4886
14.6513 | -105.4907
14.6548 | -105.4938
14.6584 | -105.4944
14.662 | -105.493
14.6655 | -105.493
14.6691 | -105.4895
14.6727 | -105.4841
14.6762 | -105.4784
14.6798 | -105.4711
14.6834 | -105.4616
14.687 | -105.4476
14.6905 | -105.4356
14.6941 | -105.4235
14.6977 | -105.4131
14.7012 | -105.3998
14.7048 | -105.384
14.7084 | -105.3657
14.7119 | -105.3505
14.7155 | -105.3381
14.7191 | -105.3262
14.7227 | -105.3126
14.7262 | -105.3008
14.7298 | -105.2922
14.7334 | -105.2855
14.7369 | -105.2786
14.7405 | -105.2733
14.7441 | -105.2693
14.7476 | -105.2661
14.7512 | -105.2626
14.7548 | -105.2596
14.7584 | -105.2609
14.7619 | -105.2659
14.7655 | -105.2702
14.7691 | -105.276
14.7726 | -105.2843
14.7762 | -105.2929
14.7798 | -105.3009
14.7833 | -105.3081
14.7869 | -105.3176
14.7905 | -105.3267
14.7941 | -105.3327
14.7976 | -105.3387
14.8012 | -105.344
14.8048 | -105.3491
14.8083 | -105.3572
14.8119 | -105.3624
14.8155 | -105.3661
14.819 | -105.3688
14.8226 | -105.3669
14.8262 | -105.3653
14.8298 | -105.362
14.8333 | -105.3558
14.8369 | -105.3483
14.8405 | -105.3421
14.844 | -105.3387
14.8476 | -105.3391
14.8512 | -105.3321
14.8547 | -105.3218
14.8583 | -105.3134
14.8619 | -105.3051
14.8655 | -105.2955
14.869 | -105.2871
14.8726 | -105.2794
14.8762 | -105.2681
14.8797 | -105.2591
14.8833 | -105.2506
14.8869 | -105.2421
14.8904 | -105.2372
14.894 | -105.2306
14.8976 | -105.2233
14.9012 | -105.2171
14.9047 | -105.2111
14.9083 | -105.2052
14.9119 | -105.2002
14.9154 | -105.1961
14.919 | -105.1957
14.9226 | -105.1965
14.9261 | -105.1959
14.9297 | -105.1948
14.9333 | -105.1985
14.9369 | -105.2017
14.9404 | -105.2054
14.944 | -105.2086
14.9476 | -105.2103
14.9511 | -105.2128
14.9547 | -105.2183
14.9583 | -105.2232
14.9618 | -105.2269
14.9654 | -105.233
14.969 | -105.2397
14.9726 | -105.2449
14.9761 | -105.2476
14.9797 | -105.2508
14.9833 | -105.2538
14.9868 | -105.256
14.9904 | -105.2577
14.994 | -105.261
14.9975 | -105.263
15.0011 | -105.2633
15.0047 | -105.265
15.0083 | -105.2694
15.0118 | -105.2714
15.0154 | -105.2713
15.019 | -105.2705
15.0225 | -105.2687
15.0261 | -105.2693
15.0297 | -105.27
15.0332 | -105.2695
15.0368 | -105.2696
15.0404 | -105.2702
15.044 | -105.2711
15.0475 | -105.2719
15.0511 | -105.2732
15.0547 | -105.2728
15.0582 | -105.2715
15.0618 | -105.2722
15.0654 | -105.2755
15.0689 | -105.2781
15.0725 | -105.2778
15.0761 | -105.2798
15.0797 | -105.2811
15.0832 | -105.2836
15.0868 | -105.2863
15.0904 | -105.2906
15.0939 | -105.2934
15.0975 | -105.2972
15.1011 | -105.3036
15.1046 | -105.308
15.1082 | -105.3126
15.1118 | -105.318
15.1154 | -105.3215
15.1189 | -105.326
15.1225 | -105.3308
15.1261 | -105.3337
15.1296 | -105.3365
15.1332 | -105.3391
15.1368 | -105.3413
15.1403 | -105.3437
15.1439 | -105.346
15.1475 | -105.346
15.1511 | -105.3471
15.1546 | -105.3481
15.1582 | -105.3504
15.1618 | -105.351
15.1653 | -105.3523
15.1689 | -105.353
15.1725 | -105.3521
15.176 | -105.3502
15.1796 | -105.3473
15.1832 | -105.3442
15.1868 | -105.3428
15.1903 | -105.3407
15.1939 | -105.3429
15.1975 | -105.3433
15.201 | -105.3437
15.2046 | -105.3443
15.2082 | -105.3457
15.2117 | -105.3478
15.2153 | -105.3522
15.2189 | -105.3577
15.2225 | -105.3646
15.226 | -105.3706
15.2296 | -105.3777
15.2332 | -105.3855
15.2367 | -105.3932
15.2403 | -105.4028
15.2439 | -105.4118
15.2474 | -105.4207
15.251 | -105.4314
15.2546 | -105.4427
15.2582 | -105.4522
15.2617 | -105.4615
15.2653 | -105.472
15.2689 | -105.4793
15.2724 | -105.4878
15.276 | -105.4961
15.2796 | -105.5009
15.2831 | -105.5044
15.2867 | -105.5056
15.2903 | -105.5087
15.2939 | -105.5087
15.2974 | -105.5081
15.301 | -105.5061
15.3046 | -105.503
15.3081 | -105.498
15.3117 | -105.4925
15.3153 | -105.4854
15.3188 | -105.4783
15.3224 | -105.4718
15.326 | -105.4641
15.3296 | -105.4571
15.3331 | -105.4497
15.3367 | -105.442
15.3403 | -105.4361
15.3438 | -105.4318
15.3474 | -105.4263
15.351 | -105.4218
15.3545 | -105.4119
15.3581 | -105.4037
15.3617 | -105.399
15.3653 | -105.3933
15.3688 | -105.3871
15.3724 | -105.3811
15.376 | -105.3737
15.3795 | -105.3694
15.3831 | -105.3666
15.3867 | -105.3635
15.3902 | -105.3556
15.3938 | -105.349
15.3974 | -105.344
15.401 | -105.3379
15.4045 | -105.3299
15.4081 | -105.3196
15.4117 | -105.3088
15.4152 | -105.3014
15.4188 | -105.2938
15.4224 | -105.2845
15.4259 | -105.2765
15.4295 | -105.2692
15.4331 | -105.2617
15.4367 | -105.2544
15.4402 | -105.2459
15.4438 | -105.2399
15.4474 | -105.2332
15.4509 | -105.2316
15.4545 | -105.2269
15.4581 | -105.2248
15.4616 | -105.222
15.4652 | -105.2206
15.4688 | -105.2206
15.4724 | -105.2216
15.4759 | -105.2224
15.4795 | -105.225
15.4831 | -105.2262
15.4866 | -105.2255
15.4902 | -105.2277
15.4938 | -105.2295
15.4973 | -105.2294
15.5009 | -105.2299
15.5045 | -105.2299
15.5081 | -105.2326
15.5116 | -105.2344
15.5152 | -105.2372
15.5188 | -105.2375
15.5223 | -105.2376
15.5259 | -105.2379
15.5295 | -105.237
15.533 | -105.2357
15.5366 | -105.2354
15.5402 | -105.2332
15.5438 | -105.2316
15.5473 | -105.2314
15.5509 | -105.2311
15.5545 | -105.2313
15.558 | -105.2281
15.5616 | -105.2229
15.5652 | -105.2193
15.5687 | -105.2158
15.5723 | -105.2109
15.5759 | -105.2087
15.5795 | -105.2059
15.583 | -105.203
15.5866 | -105.2013
15.5902 | -105.1995
15.5937 | -105.1971
15.5973 | -105.1955
15.6009 | -105.1921
15.6044 | -105.1881
15.608 | -105.1851
15.6116 | -105.1865
15.6152 | -105.1879
15.6187 | -105.1893
15.6223 | -105.1901
15.6259 | -105.1933
15.6294 | -105.1976
15.633 | -105.2002
15.6366 | -105.2037
15.6401 | -105.2076
15.6437 | -105.2131
15.6473 | -105.2162
15.6509 | -105.2194
15.6544 | -105.2252
15.658 | -105.2331
15.6616 | -105.2399
15.6651 | -105.2451
15.6687 | -105.2496
15.6723 | -105.2513
15.6758 | -105.2554
15.6794 | -105.2604
15.683 | -105.2626
15.6866 | -105.2663
15.6901 | -105.2698
15.6937 | -105.2726
15.6973 | -105.2726
15.7008 | -105.2738
15.7044 | -105.2725
15.708 | -105.27
15.7115 | -105.2687
15.7151 | -105.2659
15.7187 | -105.2627
15.7223 | -105.2624
15.7258 | -105.2582
15.7294 | -105.2531
15.733 | -105.2459
15.7365 | -105.2374
15.7401 | -105.226
15.7437 | -105.2154
15.7472 | -105.2043
15.7508 | -105.1913
15.7544 | -105.1799
15.758 | -105.1695
15.7615 | -105.1564
15.7651 | -105.1411
15.7687 | -105.1265
15.7722 | -105.1145
15.7758 | -105.1007
15.7794 | -105.0883
15.7829 | -105.0814
15.7865 | -105.0758
15.7901 | -105.0685
15.7937 | -105.0614
15.7972 | -105.0557
15.8008 | -105.051
15.8044 | -105.0482
15.8079 | -105.0467
15.8115 | -105.0475
15.8151 | -105.0485
15.8186 | -105.0511
15.8222 | -105.056
15.8258 | -105.0632
15.8294 | -105.0735
15.8329 | -105.0863
15.8365 | -105.0984
15.8401 | -105.1107
15.8436 | -105.1258
15.8472 | -105.1424
15.8508 | -105.1582
15.8543 | -105.1736
15.8579 | -105.1926
15.8615 | -105.2109
15.8651 | -105.2278
15.8686 | -105.246
15.8722 | -105.265
15.8758 | -105.2845
15.8793 | -105.3011
15.8829 | -105.3157
15.8865 | -105.3335
15.89 | -105.3499
15.8936 | -105.365
15.8972 | -105.3774
15.9008 | -105.3842
15.9043 | -105.3866
15.9079 | -105.3877
15.9115 | -105.3845
15.915 | -105.378
15.9186 | -105.3692
15.9222 | -105.3598
15.9257 | -105.3466
15.9293 | -105.3344
15.9329 | -105.321
15.9365 | -105.3044
15.94 | -105.2835
15.9436 | -105.2596
15.9472 | -105.2323
15.9507 | -105.2084
15.9543 | -105.1823
15.9579 | -105.1533
15.9614 | -105.1229
15.965 | -105.0908
15.9686 | -105.0636
15.9722 | -105.0367
15.9757 | -105.0091
15.9793 | -104.9834
15.9829 | -104.9575
15.9864 | -104.931
15.99 | -104.9052
15.9936 | -104.883
15.9971 | -104.8613
16.0007 | -104.839
16.0043 | -104.8192
16.0079 | -104.7982
16.0114 | -104.7786
16.015 | -104.7612
16.0186 | -104.7477
16.0221 | -104.7376
16.0257 | -104.7289
16.0293 | -104.7227
16.0328 | -104.7172
16.0364 | -104.7146
16.04 | -104.7139
16.0436 | -104.7138
16.0471 | -104.7157
16.0507 | -104.7189
16.0543 | -104.7245
16.0578 | -104.7319
16.0614 | -104.7403
16.065 | -104.7518
16.0685 | -104.7622
16.0721 | -104.771
16.0757 | -104.7802
16.0793 | -104.792
16.0828 | -104.8061
16.0864 | -104.8177
16.09 | -104.8297
16.0935 | -104.8424
16.0971 | -104.8551
16.1007 | -104.8671
16.1042 | -104.8766
16.1078 | -104.8853
16.1114 | -104.8934
16.115 | -104.8988
16.1185 | -104.9033
16.1221 | -104.9086
16.1257 | -104.9146
16.1292 | -104.92
16.1328 | -104.925
16.1364 | -104.9296
16.1399 | -104.9343
16.1435 | -104.9381
16.1471 | -104.9401
16.1507 | -104.9414
16.1542 | -104.9432
16.1578 | -104.9424
16.1614 | -104.9415
16.1649 | -104.939
16.1685 | -104.9395
16.1721 | -104.9423
16.1756 | -104.9429
16.1792 | -104.9402
16.1828 | -104.941
16.1864 | -104.9437
16.1899 | -104.9468
16.1935 | -104.9479
16.1971 | -104.9481
16.2006 | -104.9515
16.2042 | -104.9535
16.2078 | -104.9543
16.2113 | -104.9572
16.2149 | -104.962
16.2185 | -104.9685
16.2221 | -104.976
16.2256 | -104.9821
16.2292 | -104.9883
16.2328 | -104.9945
16.2363 | -105.0001
16.2399 | -105.0086
16.2435 | -105.0169
16.247 | -105.0227
16.2506 | -105.0279
16.2542 | -105.0338
16.2578 | -105.0393
16.2613 | -105.0447
16.2649 | -105.0492
16.2685 | -105.051
16.272 | -105.0524
16.2756 | -105.0531
16.2792 | -105.0554
16.2827 | -105.0558
16.2863 | -105.0555
16.2899 | -105.0557
16.2935 | -105.057
16.297 | -105.0561
16.3006 | -105.0585
16.3042 | -105.0624
16.3077 | -105.0645
16.3113 | -105.0637
16.3149 | -105.0623
16.3184 | -105.0637
16.322 | -105.0656
16.3256 | -105.0679
16.3292 | -105.0695
16.3327 | -105.0719
16.3363 | -105.0764
16.3399 | -105.0825
16.3434 | -105.0866
16.347 | -105.0903
16.3506 | -105.0957
16.3541 | -105.1015
16.3577 | -105.1031
16.3613 | -105.1058
16.3649 | -105.1115
16.3684 | -105.1154
16.372 | -105.1185
16.3756 | -105.1242
16.3791 | -105.1314
16.3827 | -105.1365
16.3863 | -105.1414
16.3898 | -105.1461
16.3934 | -105.1505
16.397 | -105.1569
16.4006 | -105.1618
16.4041 | -105.1662
16.4077 | -105.1754
16.4113 | -105.1857
16.4148 | -105.1937
16.4184 | -105.203
16.422 | -105.2128
16.4255 | -105.2232
16.4291 | -105.2344
16.4327 | -105.2446
16.4363 | -105.2544
16.4398 | -105.2675
16.4434 | -105.2797
16.447 | -105.2916
16.4505 | -105.3051
16.4541 | -105.319
16.4577 | -105.3351
16.4612 | -105.3521
16.4648 | -105.3702
16.4684 | -105.39
16.472 | -105.4069
16.4755 | -105.425
16.4791 | -105.4488
16.4827 | -105.4702
16.4862 | -105.4906
16.4898 | -105.5099
16.4934 | -105.5318
16.4969 | -105.5514
16.5005 | -105.5713
16.5041 | -105.5932
16.5077 | -105.6144
16.5112 | -105.6355
16.5148 | -105.653
16.5184 | -105.6708
16.5219 | -105.6877
16.5255 | -105.7026
16.5291 | -105.7164
16.5326 | -105.7305
16.5362 | -105.7401
16.5398 | -105.7467
16.5434 | -105.7572
16.5469 | -105.7654
16.5505 | -105.7719
16.5541 | -105.774
16.5576 | -105.775
16.5612 | -105.7768
16.5648 | -105.7772
16.5683 | -105.7742
16.5719 | -105.7699
16.5755 | -105.7625
16.5791 | -105.7567
16.5826 | -105.7507
16.5862 | -105.7437
16.5898 | -105.735
16.5933 | -105.726
16.5969 | -105.7158
16.6005 | -105.7037
16.604 | -105.6929
16.6076 | -105.6803
16.6112 | -105.6673
16.6148 | -105.654
16.6183 | -105.6418
16.6219 | -105.6293
16.6255 | -105.6178
16.629 | -105.6046
16.6326 | -105.5908
16.6362 | -105.5783
16.6397 | -105.5666
16.6433 | -105.5498
16.6469 | -105.537
16.6505 | -105.5273
16.654 | -105.5178
16.6576 | -105.5079
16.6612 | -105.5014
16.6647 | -105.4951
16.6683 | -105.4884
16.6719 | -105.484
16.6754 | -105.479
16.679 | -105.4741
16.6826 | -105.4713
16.6862 | -105.4685
16.6897 | -105.4671
16.6933 | -105.4657
16.6969 | -105.4665
16.7004 | -105.4696
16.704 | -105.4742
16.7076 | -105.4791
16.7111 | -105.4816
16.7147 | -105.4859
16.7183 | -105.4921
16.7219 | -105.4992
16.7254 | -105.5082
16.729 | -105.5181
16.7326 | -105.527
16.7361 | -105.5377
16.7397 | -105.5501
16.7433 | -105.5618
16.7468 | -105.5735
16.7504 | -105.5858
16.754 | -105.5996
16.7576 | -105.6143
16.7611 | -105.6271
16.7647 | -105.639
16.7683 | -105.6521
16.7718 | -105.6667
16.7754 | -105.6816
16.779 | -105.6974
16.7825 | -105.7129
16.7861 | -105.7301
16.7897 | -105.7484
16.7933 | -105.7664
16.7968 | -105.7854
16.8004 | -105.8024
16.804 | -105.8188
16.8075 | -105.8378
16.8111 | -105.857
16.8147 | -105.8788
16.8182 | -105.9
16.8218 | -105.9186
16.8254 | -105.9356
16.829 | -105.9527
16.8325 | -105.9724
16.8361 | -105.9892
16.8397 | -106.0037
16.8432 | -106.0167
16.8468 | -106.0309
16.8504 | -106.0429
16.8539 | -106.0544
16.8575 | -106.0654
16.8611 | -106.0738
16.8647 | -106.0791
16.8682 | -106.0847
16.8718 | -106.0858
16.8754 | -106.0857
16.8789 | -106.0855
16.8825 | -106.0873
16.8861 | -106.0865
16.8896 | -106.0826
16.8932 | -106.0769
16.8968 | -106.07
16.9004 | -106.0608
16.9039 | -106.0504
16.9075 | -106.0419
16.9111 | -106.0344
16.9146 | -106.0263
16.9182 | -106.0187
16.9218 | -106.0082
16.9253 | -105.9981
16.9289 | -105.9888
16.9325 | -105.9796
16.9361 | -105.9699
16.9396 | -105.9612
16.9432 | -105.9466
16.9468 | -105.934
16.9503 | -105.927
16.9539 | -105.9165
16.9575 | -105.9067
16.961 | -105.8985
16.9646 | -105.889
16.9682 | -105.8796
16.9718 | -105.8698
16.9753 | -105.8626
16.9789 | -105.8545
16.9825 | -105.8497
16.986 | -105.8464
16.9896 | -105.844
16.9932 | -105.8464
16.9967 | -105.8419
17.0003 | -105.8393
17.0039 | -105.8389
17.0075 | -105.8395
17.011 | -105.8388
17.0146 | -105.8392
17.0182 | -105.8412
17.0217 | -105.8422
17.0253 | -105.8418
17.0289 | -105.8439
17.0324 | -105.8464
17.036 | -105.8506
17.0396 | -105.853
17.0432 | -105.8561
17.0467 | -105.8604
17.0503 | -105.8636
17.0539 | -105.8653
17.0574 | -105.8647
17.061 | -105.8625
17.0646 | -105.8624
17.0681 | -105.8602
17.0717 | -105.8573
17.0753 | -105.8551
17.0789 | -105.8542
17.0824 | -105.8537
17.086 | -105.85
17.0896 | -105.8457
17.0931 | -105.8415
17.0967 | -105.835
17.1003 | -105.8279
17.1038 | -105.8215
17.1074 | -105.8161
17.111 | -105.8086
17.1146 | -105.8027
17.1181 | -105.7947
17.1217 | -105.7858
17.1253 | -105.7782
17.1288 | -105.7693
17.1324 | -105.7599
17.136 | -105.7525
17.1395 | -105.7482
17.1431 | -105.7427
17.1467 | -105.7377
17.1503 | -105.733
17.1538 | -105.7271
17.1574 | -105.7194
17.161 | -105.7127
17.1645 | -105.7078
17.1681 | -105.7022
17.1717 | -105.6959
17.1752 | -105.6891
17.1788 | -105.6831
17.1824 | -105.6752
17.186 | -105.6723
17.1895 | -105.6703
17.1931 | -105.6647
17.1967 | -105.661
17.2002 | -105.6571
17.2038 | -105.6518
17.2074 | -105.6473
17.2109 | -105.6435
17.2145 | -105.6408
17.2181 | -105.6368
17.2217 | -105.6339
17.2252 | -105.6307
17.2288 | -105.6285
17.2324 | -105.625
17.2359 | -105.6206
17.2395 | -105.6181
17.2431 | -105.6181
17.2466 | -105.6176
17.2502 | -105.617
17.2538 | -105.617
17.2574 | -105.6184
17.2609 | -105.6194
17.2645 | -105.6222
17.2681 | -105.6252
17.2716 | -105.6255
17.2752 | -105.6223
17.2788 | -105.6222
17.2823 | -105.6229
17.2859 | -105.6247
17.2895 | -105.6262
17.2931 | -105.6269
17.2966 | -105.6294
17.3002 | -105.631
17.3038 | -105.6337
17.3073 | -105.6369
17.3109 | -105.6385
17.3145 | -105.6425
17.318 | -105.6471
17.3216 | -105.6486
17.3252 | -105.6491
17.3288 | -105.6497
17.3323 | -105.6519
17.3359 | -105.6565
17.3395 | -105.6594
17.343 | -105.6619
17.3466 | -105.6659
17.3502 | -105.669
17.3537 | -105.6721
17.3573 | -105.6762
17.3609 | -105.6769
17.3645 | -105.6772
17.368 | -105.677
17.3716 | -105.6754
17.3752 | -105.6765
17.3787 | -105.678
17.3823 | -105.6788
17.3859 | -105.6796
17.3894 | -105.6791
17.393 | -105.6756
17.3966 | -105.6717
17.4002 | -105.6676
17.4037 | -105.6634
17.4073 | -105.6572
17.4109 | -105.6524
17.4144 | -105.6473
17.418 | -105.6427
17.4216 | -105.6375
17.4251 | -105.6291
17.4287 | -105.6207
17.4323 | -105.6131
17.4359 | -105.6039
17.4394 | -105.5948
17.443 | -105.5861
17.4466 | -105.5762
17.4501 | -105.5673
17.4537 | -105.5588
17.4573 | -105.5482
17.4608 | -105.5353
17.4644 | -105.5233
17.468 | -105.5124
17.4716 | -105.4992
17.4751 | -105.4851
17.4787 | -105.4735
17.4823 | -105.464
17.4858 | -105.4502
17.4894 | -105.4361
17.493 | -105.4203
17.4965 | -105.4048
17.5001 | -105.3864
17.5037 | -105.3704
17.5073 | -105.3552
17.5108 | -105.3381
17.5144 | -105.3218
17.518 | -105.3057
17.5215 | -105.2915
17.5251 | -105.2787
17.5287 | -105.263
17.5322 | -105.2478
17.5358 | -105.2335
17.5394 | -105.2201
17.543 | -105.2082
17.5465 | -105.1941
17.5501 | -105.1818
17.5537 | -105.1694
17.5572 | -105.1585
17.5608 | -105.1477
17.5644 | -105.1357
17.5679 | -105.124
17.5715 | -105.1145
17.5751 | -105.1058
17.5787 | -105.0982
17.5822 | -105.0917
17.5858 | -105.0849
17.5894 | -105.0802
17.5929 | -105.0783
17.5965 | -105.0762
17.6001 | -105.0759
17.6036 | -105.0766
17.6072 | -105.0768
17.6108 | -105.079
17.6144 | -105.0839
17.6179 | -105.0876
17.6215 | -105.0925
17.6251 | -105.0967
17.6286 | -105.1016
17.6322 | -105.1079
17.6358 | -105.1141
17.6393 | -105.1199
17.6429 | -105.129
17.6465 | -105.1368
17.6501 | -105.1458
17.6536 | -105.1565
17.6572 | -105.1659
17.6608 | -105.173
17.6643 | -105.1799
17.6679 | -105.1856
17.6715 | -105.1924
17.675 | -105.2
17.6786 | -105.2079
17.6822 | -105.2135
17.6858 | -105.2191
17.6893 | -105.2241
17.6929 | -105.2284
17.6965 | -105.2329
17.7 | -105.2353
17.7036 | -105.2393
17.7072 | -105.2434
17.7107 | -105.2473
17.7143 | -105.2491
17.7179 | -105.2512
17.7215 | -105.254
17.725 | -105.2558
17.7286 | -105.2556
17.7322 | -105.2552
17.7357 | -105.2566
17.7393 | -105.2579
17.7429 | -105.2581
17.7464 | -105.2567
17.75 | -105.2574
17.7536 | -105.2572
17.7572 | -105.256
17.7607 | -105.2555
17.7643 | -105.2521
17.7679 | -105.2494
17.7714 | -105.2482
17.775 | -105.2457
17.7786 | -105.2433
17.7821 | -105.2429
17.7857 | -105.2411
17.7893 | -105.2396
17.7929 | -105.2368
17.7964 | -105.2354
17.8 | -105.234
17.8036 | -105.2293
17.8071 | -105.2258
17.8107 | -105.2238
17.8143 | -105.2203
17.8178 | -105.2161
17.8214 | -105.214
17.825 | -105.2125
17.8286 | -105.2094
17.8321 | -105.2047
17.8357 | -105.2012
17.8393 | -105.1977
17.8428 | -105.1943
17.8464 | -105.1904
17.85 | -105.1875
17.8535 | -105.1842
17.8571 | -105.1801
17.8607 | -105.1773
17.8643 | -105.1761
17.8678 | -105.1706
17.8714 | -105.1659
17.875 | -105.1608
17.8785 | -105.1534
17.8821 | -105.1424
17.8857 | -105.1318
17.8892 | -105.1233
17.8928 | -105.1172
17.8964 | -105.1064
17.9 | -105.0957
17.9035 | -105.0834
17.9071 | -105.0704
17.9107 | -105.0586
17.9142 | -105.0465
17.9178 | -105.0353
17.9214 | -105.0253
17.9249 | -105.0124
17.9285 | -104.9996
17.9321 | -104.9879
17.9357 | -104.9765
17.9392 | -104.9633
17.9428 | -104.9506
17.9464 | -104.9403
17.9499 | -104.9289
17.9535 | -104.9182
17.9571 | -104.9076
17.9606 | -104.9007
17.9642 | -104.8937
17.9678 | -104.8858
17.9714 | -104.8769
17.9749 | -104.8702
17.9785 | -104.8668
17.9821 | -104.8624
17.9856 | -104.8568
17.9892 | -104.8526
17.9928 | -104.8509
17.9963 | -104.8505
17.9999 | -104.8517
18.0035 | -104.8542
18.0071 | -104.8554
18.0106 | -104.8566
18.0142 | -104.859
18.0178 | -104.8646
18.0213 | -104.8694
18.0249 | -104.8745
18.0285 | -104.8782
18.032 | -104.8821
18.0356 | -104.8853
18.0392 | -104.8881
18.0428 | -104.8904
18.0463 | -104.8916
18.0499 | -104.8936
18.0535 | -104.8929
18.057 | -104.8938
18.0606 | -104.8937
18.0642 | -104.8916
18.0677 | -104.8863
18.0713 | -104.8814
18.0749 | -104.878
18.0785 | -104.8704
18.082 | -104.8614
18.0856 | -104.8541
18.0892 | -104.8463
18.0927 | -104.835
18.0963 | -104.8198
18.0999 | -104.8024
18.1034 | -104.7855
18.107 | -104.7683
18.1106 | -104.7513
18.1142 | -104.7335
18.1177 | -104.7152
18.1213 | -104.6946
18.1249 | -104.6768
18.1284 | -104.6578
18.132 | -104.6388
18.1356 | -104.6216
18.1391 | -104.6048
18.1427 | -104.5851
18.1463 | -104.5614
18.1499 | -104.5379
18.1534 | -104.5156
18.157 | -104.4902
18.1606 | -104.4651
18.1641 | -104.4384
18.1677 | -104.4112
18.1713 | -104.3865
18.1748 | -104.3587
18.1784 | -104.3284
18.182 | -104.2994
18.1855 | -104.2675
18.1891 | -104.2361
18.1927 | -104.2059
18.1963 | -104.1784
18.1998 | -104.151
18.2034 | -104.1221
18.207 | -104.0931
18.2105 | -104.0639
18.2141 | -104.0341
18.2177 | -104.0036
18.2213 | -103.9731
18.2248 | -103.9458
18.2284 | -103.9171
18.232 | -103.8877
18.2355 | -103.8583
18.2391 | -103.8323
18.2427 | -103.8078
18.2462 | -103.7842
18.2498 | -103.7597
18.2534 | -103.7369
18.2569 | -103.7122
18.2605 | -103.6889
18.2641 | -103.6671
18.2677 | -103.6466
18.2712 | -103.6279
18.2748 | -103.6087
18.2784 | -103.5898
18.2819 | -103.5732
18.2855 | -103.5586
18.2891 | -103.5463
18.2926 | -103.5329
18.2962 | -103.5201
18.2998 | -103.5087
18.3034 | -103.4988
18.3069 | -103.4905
18.3105 | -103.4826
18.3141 | -103.4751
18.3176 | -103.4677
18.3212 | -103.4627
18.3248 | -103.4573
18.3284 | -103.4517
18.3319 | -103.4485
18.3355 | -103.4449
18.3391 | -103.4426
18.3426 | -103.4398
18.3462 | -103.4372
18.3498 | -103.4356
18.3533 | -103.4353
18.3569 | -103.4358
18.3605 | -103.4365
18.364 | -103.4367
18.3676 | -103.4347
18.3712 | -103.4314
18.3748 | -103.428
18.3783 | -103.4249
18.3819 | -103.4228
18.3855 | -103.4191
18.389 | -103.4156
18.3926 | -103.4093
18.3962 | -103.4023
18.3998 | -103.3972
18.4033 | -103.3912
18.4069 | -103.3852
18.4105 | -103.3784
18.414 | -103.3724
18.4176 | -103.366
18.4212 | -103.3605
18.4247 | -103.3565
18.4283 | -103.3504
18.4319 | -103.3421
18.4354 | -103.337
18.439 | -103.3338
18.4426 | -103.3311
18.4462 | -103.3283
18.4497 | -103.3242
18.4533 | -103.3198
18.4569 | -103.3168
18.4604 | -103.3142
18.464 | -103.3113
18.4676 | -103.3091
18.4712 | -103.3092
18.4747 | -103.3086
18.4783 | -103.3074
18.4819 | -103.3086
18.4854 | -103.3076
18.489 | -103.3086
18.4926 | -103.3098
18.4961 | -103.3099
18.4997 | -103.3092
18.5033 | -103.3073
18.5069 | -103.3044
18.5104 | -103.303
18.514 | -103.3026
18.5176 | -103.305
18.5211 | -103.3067
18.5247 | -103.3074
18.5283 | -103.309
18.5318 | -103.3117
18.5354 | -103.3148
18.539 | -103.321
18.5425 | -103.3268
18.5461 | -103.3318
18.5497 | -103.3373
18.5533 | -103.3425
18.5568 | -103.351
18.5604 | -103.3606
18.564 | -103.3708
18.5675 | -103.3796
18.5711 | -103.3868
18.5747 | -103.3959
18.5783 | -103.4069
18.5818 | -103.416
18.5854 | -103.4236
18.589 | -103.4321
18.5925 | -103.4423
18.5961 | -103.4498
18.5997 | -103.4574
18.6032 | -103.4664
18.6068 | -103.4743
18.6104 | -103.4812
18.6139 | -103.4876
18.6175 | -103.4948
18.6211 | -103.5024
18.6247 | -103.5089
18.6282 | -103.5135
18.6318 | -103.5199
18.6354 | -103.5265
18.6389 | -103.534
18.6425 | -103.5387
18.6461 | -103.5409
18.6497 | -103.5445
18.6532 | -103.549
18.6568 | -103.5532
18.6604 | -103.5563
18.6639 | -103.5591
18.6675 | -103.5616
18.6711 | -103.5633
18.6746 | -103.5646
18.6782 | -103.5649
18.6818 | -103.566
18.6853 | -103.5674
18.6889 | -103.5681
18.6925 | -103.5711
18.6961 | -103.5762
18.6996 | -103.5805
18.7032 | -103.5846
18.7068 | -103.5891
18.7103 | -103.5945
18.7139 | -103.6013
18.7175 | -103.6084
18.7211 | -103.6162
18.7246 | -103.6254
18.7282 | -103.6363
18.7318 | -103.6479
18.7353 | -103.6605
18.7389 | -103.6765
18.7425 | -103.6932
18.746 | -103.7098
18.7496 | -103.7298
18.7532 | -103.7502
18.7568 | -103.7708
18.7603 | -103.7923
18.7639 | -103.8153
18.7675 | -103.8399
18.771 | -103.862
18.7746 | -103.8835
18.7782 | -103.905
18.7817 | -103.9268
18.7853 | -103.9477
18.7889 | -103.9685
18.7924 | -103.9888
18.796 | -104.008
18.7996 | -104.0266
18.8032 | -104.0448
18.8067 | -104.0614
18.8103 | -104.0757
18.8139 | -104.088
18.8174 | -104.0966
18.821 | -104.1027
18.8246 | -104.1086
18.8281 | -104.1113
18.8317 | -104.1109
18.8353 | -104.1128
18.8389 | -104.1141
18.8424 | -104.1106
18.846 | -104.1037
18.8496 | -104.0961
18.8531 | -104.0877
18.8567 | -104.0745
18.8603 | -104.0574
18.8638 | -104.0394
18.8674 | -104.0204
18.871 | -103.9997
18.8746 | -103.9788
18.8781 | -103.9561
18.8817 | -103.9331
18.8853 | -103.9077
18.8888 | -103.8809
18.8924 | -103.8535
18.896 | -103.8265
18.8995 | -103.7971
18.9031 | -103.7664
18.9067 | -103.7358
18.9103 | -103.7084
18.9138 | -103.6815
18.9174 | -103.654
18.921 | -103.628
18.9245 | -103.6052
18.9281 | -103.5827
18.9317 | -103.5609
18.9352 | -103.5407
18.9388 | -103.521
18.9424 | -103.5034
18.946 | -103.4888
18.9495 | -103.4752
18.9531 | -103.4613
18.9567 | -103.4496
18.9602 | -103.4399
18.9638 | -103.4304
18.9674 | -103.4237
18.9709 | -103.42
18.9745 | -103.417
18.9781 | -103.4172
18.9817 | -103.4191
18.9852 | -103.423
18.9888 | -103.4288
18.9924 | -103.4376
18.9959 | -103.4492
18.9995 | -103.4622
19.0031 | -103.476
19.0066 | -103.4904
19.0102 | -103.5056
19.0138 | -103.5236
19.0174 | -103.5415
19.0209 | -103.5601
19.0245 | -103.581
19.0281 | -103.6023
19.0316 | -103.6238
19.0352 | -103.6467
19.0388 | -103.6701
19.0423 | -103.6925
19.0459 | -103.7145
19.0495 | -103.7353
19.0531 | -103.7448
19.0566 | -103.7607
19.0602 | -103.7761
19.0638 | -103.7891
19.0673 | -103.7983
19.0709 | -103.8067
19.0745 | -103.8143
19.078 | -103.821
19.0816 | -103.8252
19.0852 | -103.8276
19.0888 | -103.8292
19.0923 | -103.8283
19.0959 | -103.822
19.0995 | -103.8146
19.103 | -103.8033
19.1066 | -103.7882
19.1102 | -103.7706
19.1137 | -103.7527
19.1173 | -103.732
19.1209 | -103.7091
19.1245 | -103.6853
19.128 | -103.6617
19.1316 | -103.6357
19.1352 | -103.6071
19.1387 | -103.577
19.1423 | -103.546
19.1459 | -103.5138
19.1494 | -103.4827
19.153 | -103.4515
19.1566 | -103.4221
19.1602 | -103.3923
19.1637 | -103.3636
19.1673 | -103.3364
19.1709 | -103.3114
19.1744 | -103.2884
19.178 | -103.2651
19.1816 | -103.2433
19.1851 | -103.2247
19.1887 | -103.207
19.1923 | -103.1886
19.1959 | -103.175
19.1994 | -103.1656
19.203 | -103.1586
19.2066 | -103.155
19.2101 | -103.1556
19.2137 | -103.1574
19.2173 | -103.1615
19.2208 | -103.1698
19.2244 | -103.1799
19.228 | -103.1919
19.2316 | -103.2071
19.2351 | -103.2262
19.2387 | -103.2465
19.2423 | -103.2686
19.2458 | -103.2938
19.2494 | -103.3196
19.253 | -103.3459
19.2565 | -103.3755
19.2601 | -103.4054
19.2637 | -103.4345
19.2673 | -103.4658
19.2708 | -103.4998
19.2744 | -103.5314
19.278 | -103.5629
19.2815 | -103.5954
19.2851 | -103.6262
19.2887 | -103.6583
19.2922 | -103.6909
19.2958 | -103.7206
19.2994 | -103.7502
19.303 | -103.7786
19.3065 | -103.8031
19.3101 | -103.8271
19.3137 | -103.8504
19.3172 | -103.8703
19.3208 | -103.8868
19.3244 | -103.906
19.3279 | -103.9241
19.3315 | -103.9394
19.3351 | -103.9528
19.3387 | -103.9644
19.3422 | -103.9758
19.3458 | -103.985
19.3494 | -103.9905
19.3529 | -103.9945
19.3565 | -103.9988
19.3601 | -104.0024
19.3636 | -104.0045
19.3672 | -104.0087
19.3708 | -104.015
19.3744 | -104.0205
19.3779 | -104.0272
19.3815 | -104.0343
19.3851 | -104.0419
19.3886 | -104.0489
19.3922 | -104.056
19.3958 | -104.0654
19.3993 | -104.0737
19.4029 | -104.0821
19.4065 | -104.0937
19.4101 | -104.1069
19.4136 | -104.1201
19.4172 | -104.1341
19.4208 | -104.1512
19.4243 | -104.1661
19.4279 | -104.1819
19.4315 | -104.1976
19.435 | -104.2138
19.4386 | -104.2313
19.4422 | -104.2487
19.4458 | -104.2662
19.4493 | -104.2838
19.4529 | -104.3013
19.4565 | -104.3188
19.46 | -104.3369
19.4636 | -104.3557
19.4672 | -104.3757
19.4707 | -104.3926
19.4743 | -104.4086
19.4779 | -104.4209
19.4815 | -104.4333
19.485 | -104.446
19.4886 | -104.4593
19.4922 | -104.4718
19.4957 | -104.4834
19.4993 | -104.4953
19.5029 | -104.5058
19.5064 | -104.5146
19.51 | -104.5216
19.5136 | -104.5288
19.5172 | -104.5343
19.5207 | -104.5377
19.5243 | -104.5405
19.5279 | -104.542
19.5314 | -104.5401
19.535 | -104.5372
19.5386 | -104.5366
19.5421 | -104.5332
19.5457 | -104.5292
19.5493 | -104.5254
19.5529 | -104.5202
19.5564 | -104.5143
19.56 | -104.5075
19.5636 | -104.4987
19.5671 | -104.4886
19.5707 | -104.4781
19.5743 | -104.4703
19.5778 | -104.4615
19.5814 | -104.4519
19.585 | -104.4426
19.5886 | -104.4362
19.5921 | -104.431
19.5957 | -104.4261
19.5993 | -104.4227
19.6028 | -104.4175
19.6064 | -104.4124
19.61 | -104.4083
19.6135 | -104.404
19.6171 | -104.4026
19.6207 | -104.4015
19.6243 | -104.4016
19.6278 | -104.4052
19.6314 | -104.4093
19.635 | -104.4134
19.6385 | -104.419
19.6421 | -104.4267
19.6457 | -104.434
19.6492 | -104.4401
19.6528 | -104.4466
19.6564 | -104.4547
19.66 | -104.4639
19.6635 | -104.4753
19.6671 | -104.485
19.6707 | -104.4964
19.6742 | -104.5083
19.6778 | -104.5213
19.6814 | -104.5335
19.6849 | -104.5436
19.6885 | -104.5543
19.6921 | -104.5664
19.6957 | -104.5804
19.6992 | -104.593
19.7028 | -104.6081
19.7064 | -104.6204
19.7099 | -104.6301
19.7135 | -104.6419
19.7171 | -104.6514
19.7206 | -104.6575
19.7242 | -104.6622
19.7278 | -104.6676
19.7314 | -104.6741
19.7349 | -104.6779
19.7385 | -104.6801
19.7421 | -104.6807
19.7456 | -104.6809
19.7492 | -104.6797
19.7528 | -104.678
19.7563 | -104.6779
19.7599 | -104.6754
19.7635 | -104.6709
19.7671 | -104.663
19.7706 | -104.6539
19.7742 | -104.641
19.7778 | -104.6257
19.7813 | -104.6134
19.7849 | -104.6004
19.7885 | -104.5869
19.792 | -104.5718
19.7956 | -104.555
19.7992 | -104.5409
19.8028 | -104.5285
19.8063 | -104.5141
19.8099 | -104.4984
19.8135 | -104.4834
19.817 | -104.4673
19.8206 | -104.4516
19.8242 | -104.4368
19.8277 | -104.4242
19.8313 | -104.4128
19.8349 | -104.4024
19.8385 | -104.3949
19.842 | -104.3871
19.8456 | -104.3791
19.8492 | -104.3752
19.8527 | -104.374
19.8563 | -104.3728
19.8599 | -104.3734
19.8634 | -104.3762
19.867 | -104.3802
19.8706 | -104.3835
19.8742 | -104.3876
19.8777 | -104.3953
19.8813 | -104.4052
19.8849 | -104.4154
19.8884 | -104.4282
19.892 | -104.4428
19.8956 | -104.4594
19.8991 | -104.475
19.9027 | -104.4898
19.9063 | -104.5045
19.9099 | -104.5192
19.9134 | -104.5356
19.917 | -104.5526
19.9206 | -104.5706
19.9241 | -104.5877
19.9277 | -104.603
19.9313 | -104.6185
19.9348 | -104.6346
19.9384 | -104.6495
19.942 | -104.6635
19.9456 | -104.6731
19.9491 | -104.685
19.9527 | -104.696
19.9563 | -104.7049
19.9598 | -104.7114
19.9634 | -104.7152
19.967 | -104.7165
19.9705 | -104.7178
19.9741 | -104.7147
19.9777 | -104.7082
19.9813 | -104.7016
19.9848 | -104.6928
19.9884 | -104.6844
19.992 | -104.6765
19.9955 | -104.6656
19.9991 | -104.6535
20.0027 | -104.6423
20.0062 | -104.6277
20.0098 | -104.6118
20.0134 | -104.5938
20.017 | -104.5748
20.0205 | -104.5541
20.0241 | -104.5344
20.0277 | -104.5149
20.0312 | -104.4953
20.0348 | -104.4744
20.0384 | -104.4554
20.0419 | -104.4381
20.0455 | -104.4203
20.0491 | -104.4056
20.0527 | -104.3911
20.0562 | -104.3789
20.0598 | -104.3686
20.0634 | -104.3569
20.0669 | -104.347
20.0705 | -104.337
20.0741 | -104.3314
20.0776 | -104.3289
20.0812 | -104.3252
20.0848 | -104.3368
20.0884 | -104.3348
20.0919 | -104.336
20.0955 | -104.3402
20.0991 | -104.3445
20.1026 | -104.3498
20.1062 | -104.3575
20.1098 | -104.3663
20.1133 | -104.377
20.1169 | -104.3896
20.1205 | -104.4017
20.1241 | -104.4156
20.1276 | -104.4299
20.1312 | -104.4461
20.1348 | -104.4615
20.1383 | -104.4759
20.1419 | -104.4927
20.1455 | -104.5102
20.149 | -104.524
20.1526 | -104.5409
20.1562 | -104.559
20.1598 | -104.5747
20.1633 | -104.5913
20.1669 | -104.6061
20.1705 | -104.6204
20.174 | -104.6344
20.1776 | -104.6475
20.1812 | -104.6597
20.1847 | -104.6732
20.1883 | -104.6882
20.1919 | -104.702
20.1955 | -104.7127
20.199 | -104.7227
20.2026 | -104.7319
20.2062 | -104.7389
20.2097 | -104.7441
20.2133 | -104.749
20.2169 | -104.7524
20.2204 | -104.7528
20.224 | -104.7544
20.2276 | -104.7562
20.2312 | -104.7592
20.2347 | -104.7584
20.2383 | -104.7565
20.2419 | -104.7535
20.2454 | -104.7507
20.249 | -104.7497
20.2526 | -104.7494
20.2561 | -104.7463
20.2597 | -104.7427
20.2633 | -104.7393
20.2669 | -104.7339
20.2704 | -104.7288
20.274 | -104.7232
20.2776 | -104.7171
20.2811 | -104.7131
20.2847 | -104.709
20.2883 | -104.7039
20.2918 | -104.6997
20.2954 | -104.6949
20.299 | -104.6915
20.3026 | -104.6873
20.3061 | -104.6823
20.3097 | -104.6766
20.3133 | -104.6703
20.3168 | -104.6651
20.3204 | -104.6606
20.324 | -104.6541
20.3275 | -104.6457
20.3311 | -104.6376
20.3347 | -104.6311
20.3383 | -104.6238
20.3418 | -104.6169
20.3454 | -104.6108
20.349 | -104.6047
20.3525 | -104.5973
20.3561 | -104.5903
20.3597 | -104.5803
20.3632 | -104.5712
20.3668 | -104.5617
20.3704 | -104.5513
20.374 | -104.5393
20.3775 | -104.5268
20.3811 | -104.5154
20.3847 | -104.5031
20.3882 | -104.4906
20.3918 | -104.4764
20.3954 | -104.4628
20.3989 | -104.4496
20.4025 | -104.436
20.4061 | -104.4207
20.4097 | -104.4074
20.4132 | -104.3927
20.4168 | -104.3777
20.4204 | -104.3637
20.4239 | -104.35
20.4275 | -104.3381
20.4311 | -104.3268
20.4346 | -104.3148
20.4382 | -104.3026
20.4418 | -104.2871
20.4454 | -104.2735
20.4489 | -104.2627
20.4525 | -104.2503
20.4561 | -104.2407
20.4596 | -104.2317
20.4632 | -104.2221
20.4668 | -104.2115
20.4703 | -104.2014
20.4739 | -104.1918
20.4775 | -104.1813
20.4811 | -104.1734
20.4846 | -104.1666
20.4882 | -104.1572
20.4918 | -104.1472
20.4953 | -104.1393
20.4989 | -104.1318
20.5025 | -104.1244
20.506 | -104.1174
20.5096 | -104.1111
20.5132 | -104.1025
20.5168 | -104.0927
20.5203 | -104.084
20.5239 | -104.076
20.5275 | -104.0665
20.531 | -104.0564
20.5346 | -104.0479
20.5382 | -104.0393
20.5417 | -104.0307
20.5453 | -104.0213
20.5489 | -104.0121
20.5525 | -104.0027
20.556 | -103.9943
20.5596 | -103.9847
20.5632 | -103.9719
20.5667 | -103.9606
20.5703 | -103.9493
20.5739 | -103.9378
20.5774 | -103.9267
20.581 | -103.9165
20.5846 | -103.9054
20.5882 | -103.8929
20.5917 | -103.8815
20.5953 | -103.8685
20.5989 | -103.8541
20.6024 | -103.8432
20.606 | -103.8308
20.6096 | -103.8191
20.6131 | -103.8085
20.6167 | -103.7972
20.6203 | -103.787
20.6239 | -103.7765
20.6274 | -103.7642
20.631 | -103.7521
20.6346 | -103.7402
20.6381 | -103.7274
20.6417 | -103.7137
20.6453 | -103.7007
20.6488 | -103.6878
20.6524 | -103.6758
20.656 | -103.6645
20.6596 | -103.6524
20.6631 | -103.6394
20.6667 | -103.628
20.6703 | -103.6197
20.6738 | -103.6112
20.6774 | -103.6026
20.681 | -103.5948
20.6845 | -103.5864
20.6881 | -103.5743
20.6917 | -103.5598
20.6953 | -103.5482
20.6988 | -103.5383
20.7024 | -103.53
20.706 | -103.5226
20.7095 | -103.5162
20.7131 | -103.5083
20.7167 | -103.4999
20.7202 | -103.49
20.7238 | -103.4802
20.7274 | -103.4723
20.731 | -103.4679
20.7345 | -103.4636
20.7381 | -103.455
20.7417 | -103.4446
20.7452 | -103.4361
20.7488 | -103.4267
20.7524 | -103.4138
20.7559 | -103.4023
20.7595 | -103.3908
20.7631 | -103.3751
20.7667 | -103.3602
20.7702 | -103.3469
20.7738 | -103.3326
20.7774 | -103.319
20.7809 | -103.3058
20.7845 | -103.2941
20.7881 | -103.2798
20.7916 | -103.2647
20.7952 | -103.2496
20.7988 | -103.2342
20.8024 | -103.2179
20.8059 | -103.2021
20.8095 | -103.1876
20.8131 | -103.1721
20.8166 | -103.1574
20.8202 | -103.1447
20.8238 | -103.1317
20.8273 | -103.1176
20.8309 | -103.1037
20.8345 | -103.0898
20.8381 | -103.0785
20.8416 | -103.0679
20.8452 | -103.0574
20.8488 | -103.0491
20.8523 | -103.0436
20.8559 | -103.0406
20.8595 | -103.0361
20.863 | -103.0319
20.8666 | -103.0288
20.8702 | -103.027
20.8738 | -103.027
20.8773 | -103.0294
20.8809 | -103.033
20.8845 | -103.0377
20.888 | -103.0459
20.8916 | -103.0542
20.8952 | -103.0609
20.8987 | -103.0693
20.9023 | -103.0804
20.9059 | -103.0927
20.9095 | -103.1043
20.913 | -103.119
20.9166 | -103.1353
20.9202 | -103.1518
20.9237 | -103.1697
20.9273 | -103.1891
20.9309 | -103.21
20.9344 | -103.2333
20.938 | -103.2572
20.9416 | -103.2802
20.9452 | -103.3016
20.9487 | -103.3206
20.9523 | -103.3392
20.9559 | -103.36
20.9594 | -103.3812
20.963 | -103.4047
20.9666 | -103.4243
20.9701 | -103.44
20.9737 | -103.4545
20.9773 | -103.4716
20.9809 | -103.4905
20.9844 | -103.5096
20.988 | -103.5273
20.9916 | -103.5377
20.9951 | -103.5443
20.9987 | -103.5487
21.0023 | -103.5519
21.0058 | -103.5545
21.0094 | -103.5572
21.013 | -103.5588
21.0166 | -103.5596
21.0201 | -103.5591
21.0237 | -103.5558
21.0273 | -103.5501
21.0308 | -103.5433
21.0344 | -103.5369
21.038 | -103.5282
21.0415 | -103.5175
21.0451 | -103.5063
21.0487 | -103.4976
21.0523 | -103.4867
21.0558 | -103.477
21.0594 | -103.4707
21.063 | -103.4595
21.0665 | -103.4466
21.0701 | -103.4339
21.0737 | -103.4204
21.0772 | -103.4066
21.0808 | -103.3932
21.0844 | -103.3818
21.088 | -103.3714
21.0915 | -103.3623
21.0951 | -103.3537
21.0987 | -103.3469
21.1022 | -103.3394
21.1058 | -103.3302
21.1094 | -103.3248
21.1129 | -103.3212
21.1165 | -103.3236
21.1201 | -103.3201
21.1237 | -103.318
21.1272 | -103.3164
21.1308 | -103.316
21.1344 | -103.3143
21.1379 | -103.3114
21.1415 | -103.3096
21.1451 | -103.3098
21.1486 | -103.3093
21.1522 | -103.3097
21.1558 | -103.3123
21.1594 | -103.3158
21.1629 | -103.3176
21.1665 | -103.3187
21.1701 | -103.3197
21.1736 | -103.3209
21.1772 | -103.3205
21.1808 | -103.3223
21.1843 | -103.3223
21.1879 | -103.3222
21.1915 | -103.3215
21.1951 | -103.3209
21.1986 | -103.3195
21.2022 | -103.3149
21.2058 | -103.3075
21.2093 | -103.3024
21.2129 | -103.2965
21.2165 | -103.2904
21.22 | -103.2843
21.2236 | -103.2779
21.2272 | -103.2701
21.2308 | -103.2601
21.2343 | -103.2502
21.2379 | -103.2413
21.2415 | -103.231
21.245 | -103.2212
21.2486 | -103.2155
21.2522 | -103.209
21.2557 | -103.1999
21.2593 | -103.1906
21.2629 | -103.1797
21.2665 | -103.1693
21.27 | -103.1611
21.2736 | -103.1517
21.2772 | -103.1361
21.2807 | -103.1217
21.2843 | -103.111
21.2879 | -103.1016
21.2914 | -103.0931
21.295 | -103.0844
21.2986 | -103.0745
21.3022 | -103.0644
21.3057 | -103.0543
21.3093 | -103.043
21.3129 | -103.0321
21.3164 | -103.0204
21.32 | -103.0076
21.3236 | -102.9954
21.3271 | -102.9832
21.3307 | -102.9705
21.3343 | -102.958
21.3379 | -102.9463
21.3414 | -102.9337
21.345 | -102.9209
21.3486 | -102.9067
21.3521 | -102.8932
21.3557 | -102.8801
21.3593 | -102.8667
21.3628 | -102.8535
21.3664 | -102.8402
21.37 | -102.827
21.3736 | -102.8143
21.3771 | -102.8015
21.3807 | -102.7892
21.3843 | -102.7757
21.3878 | -102.7614
21.3914 | -102.7464
21.395 | -102.7333
21.3985 | -102.7197
21.4021 | -102.7042
21.4057 | -102.6908
21.4093 | -102.6763
21.4128 | -102.6628
21.4164 | -102.6515
21.42 | -102.6404
21.4235 | -102.6294
21.4271 | -102.6202
21.4307 | -102.6096
21.4342 | -102.6005
21.4378 | -102.593
21.4414 | -102.5864
21.445 | -102.5779
21.4485 | -102.5666
21.4521 | -102.5568
21.4557 | -102.5479
21.4592 | -102.5386
21.4628 | -102.5329
21.4664 | -102.5294
21.4699 | -102.5271
21.4735 | -102.5249
21.4771 | -102.5205
21.4807 | -102.5174
21.4842 | -102.5131
21.4878 | -102.512
21.4914 | -102.5096
21.4949 | -102.5089
21.4985 | -102.5079
21.5021 | -102.5072
21.5056 | -102.5081
21.5092 | -102.509
21.5128 | -102.5065
21.5164 | -102.5054
21.5199 | -102.5029
21.5235 | -102.5005
21.5271 | -102.4974
21.5306 | -102.4966
21.5342 | -102.4954
21.5378 | -102.4936
21.5413 | -102.4903
21.5449 | -102.4881
21.5485 | -102.4868
21.5521 | -102.4827
21.5556 | -102.477
21.5592 | -102.4717
21.5628 | -102.4665
21.5663 | -102.4618
21.5699 | -102.4582
21.5735 | -102.4553
21.577 | -102.4536
21.5806 | -102.4561
21.5842 | -102.4551
21.5878 | -102.4507
21.5913 | -102.446
21.5949 | -102.4425
21.5985 | -102.441
21.602 | -102.438
21.6056 | -102.4332
21.6092 | -102.4288
21.6127 | -102.4264
21.6163 | -102.4217
21.6199 | -102.4174
21.6235 | -102.4154
21.627 | -102.4137
21.6306 | -102.4082
21.6342 | -102.4064
21.6377 | -102.406
21.6413 | -102.4037
21.6449 | -102.401
21.6484 | -102.3988
21.652 | -102.397
21.6556 | -102.3971
21.6592 | -102.3962
21.6627 | -102.3934
21.6663 | -102.3901
21.6699 | -102.3888
21.6734 | -102.3869
21.677 | -102.3821
21.6806 | -102.3778
21.6841 | -102.3762
21.6877 | -102.3732
21.6913 | -102.3703
21.6949 | -102.3675
21.6984 | -102.3653
21.702 | -102.3623
21.7056 | -102.3586
21.7091 | -102.3561
21.7127 | -102.3522
21.7163 | -102.3492
21.7198 | -102.3453
21.7234 | -102.3417
21.727 | -102.3401
21.7306 | -102.337
21.7341 | -102.3325
21.7377 | -102.3308
21.7413 | -102.328
21.7448 | -102.3255
21.7484 | -102.3227
21.752 | -102.3214
21.7555 | -102.3212
21.7591 | -102.3184
21.7627 | -102.3134
21.7663 | -102.3089
21.7698 | -102.3027
21.7734 | -102.2972
21.777 | -102.2925
21.7805 | -102.2876
21.7841 | -102.2836
21.7877 | -102.2771
21.7912 | -102.2729
21.7948 | -102.268
21.7984 | -102.2625
21.802 | -102.2547
21.8055 | -102.247
21.8091 | -102.2435
21.8127 | -102.242
21.8162 | -102.2405
21.8198 | -102.2381
21.8234 | -102.2335
21.8269 | -102.2299
21.8305 | -102.2293
21.8341 | -102.2302
21.8377 | -102.2301
21.8412 | -102.2322
21.8448 | -102.2335
21.8484 | -102.2363
21.8519 | -102.2412
21.8555 | -102.2462
21.8591 | -102.2515
21.8626 | -102.2558
21.8662 | -102.2622
21.8698 | -102.2689
21.8734 | -102.2729
21.8769 | -102.2756
21.8805 | -102.2794
21.8841 | -102.2833
21.8876 | -102.285
21.8912 | -102.2887
21.8948 | -102.2933
21.8983 | -102.2992
21.9019 | -102.3051
21.9055 | -102.3113
21.9091 | -102.3173
21.9126 | -102.3228
21.9162 | -102.3289
21.9198 | -102.3334
21.9233 | -102.3382
21.9269 | -102.3421
21.9305 | -102.3442
21.934 | -102.344
21.9376 | -102.346
21.9412 | -102.347
21.9448 | -102.3464
21.9483 | -102.3447
21.9519 | -102.3407
21.9555 | -102.3361
21.959 | -102.3308
21.9626 | -102.3263
21.9662 | -102.319
21.9697 | -102.3112
21.9733 | -102.3044
21.9769 | -102.2966
21.9805 | -102.2883
21.984 | -102.2804
21.9876 | -102.2734
21.9912 | -102.2675
21.9947 | -102.2597
21.9983 | -102.2508
22.0019 | -102.2424
22.0054 | -102.2347
22.009 | -102.2256
22.0126 | -102.2142
22.0162 | -102.204
22.0197 | -102.1953
22.0233 | -102.1876
22.0269 | -102.1795
22.0304 | -102.1708
22.034 | -102.1617
22.0376 | -102.1533
22.0411 | -102.1466
22.0447 | -102.1419
22.0483 | -102.1372
22.0519 | -102.1306
22.0554 | -102.1264
22.059 | -102.1226
22.0626 | -102.1206
22.0661 | -102.12
22.0697 | -102.121
22.0733 | -102.1238
22.0768 | -102.1268
22.0804 | -102.1303
22.084 | -102.1337
22.0876 | -102.1341
22.0911 | -102.1351
22.0947 | -102.1417
22.0983 | -102.1518
22.1018 | -102.1603
22.1054 | -102.1688
22.109 | -102.1784
22.1125 | -102.1897
22.1161 | -102.204
22.1197 | -102.2152
22.1233 | -102.2247
22.1268 | -102.2366
22.1304 | -102.2525
22.134 | -102.2699
22.1375 | -102.2883
22.1411 | -102.3033
22.1447 | -102.3204
22.1482 | -102.3355
22.1518 | -102.3524
22.1554 | -102.3684
22.159 | -102.3844
22.1625 | -102.4017
22.1661 | -102.4195
22.1697 | -102.4364
22.1732 | -102.4543
22.1768 | -102.4752
22.1804 | -102.4935
22.1839 | -102.5115
22.1875 | -102.5309
22.1911 | -102.551
22.1947 | -102.5715
22.1982 | -102.5932
22.2018 | -102.6146
22.2054 | -102.6349
22.2089 | -102.6546
22.2125 | -102.6729
22.2161 | -102.6913
22.2196 | -102.7066
22.2232 | -102.7244
22.2268 | -102.743
22.2304 | -102.7585
22.2339 | -102.7756
22.2375 | -102.7929
22.2411 | -102.8097
22.2446 | -102.8273
22.2482 | -102.8472
22.2518 | -102.866
22.2553 | -102.884
22.2589 | -102.8991
22.2625 | -102.9125
22.2661 | -102.9304
22.2696 | -102.9501
22.2732 | -102.969
22.2768 | -102.9836
22.2803 | -102.9988
22.2839 | -103.0145
22.2875 | -103.0297
22.291 | -103.0451
22.2946 | -103.0592
22.2982 | -103.0738
22.3018 | -103.0879
22.3053 | -103.0982
22.3089 | -103.1077
22.3125 | -103.1192
22.316 | -103.1292
22.3196 | -103.1382
22.3232 | -103.1494
22.3267 | -103.1616
22.3303 | -103.1717
22.3339 | -103.1797
22.3375 | -103.1869
22.341 | -103.1922
22.3446 | -103.1996
22.3482 | -103.2068
22.3517 | -103.2125
22.3553 | -103.2194
22.3589 | -103.2258
22.3624 | -103.2313
22.366 | -103.2407
22.3696 | -103.2484
22.3732 | -103.2573
22.3767 | -103.2658
22.3803 | -103.2732
22.3839 | -103.2818
22.3874 | -103.2908
22.391 | -103.3017
22.3946 | -103.3134
22.3981 | -103.3267
22.4017 | -103.3453
22.4053 | -103.363
22.4089 | -103.375
22.4124 | -103.3869
22.416 | -103.4011
22.4196 | -103.4169
22.4231 | -103.4347
22.4267 | -103.452
22.4303 | -103.4671
22.4338 | -103.4847
22.4374 | -103.5045
22.441 | -103.524
22.4446 | -103.5419
22.4481 | -103.5587
22.4517 | -103.5762
22.4553 | -103.5965
22.4588 | -103.6165
22.4624 | -103.6341
22.466 | -103.6516
22.4695 | -103.6664
22.4731 | -103.6791
22.4767 | -103.6938
22.4803 | -103.7056
22.4838 | -103.7146
22.4874 | -103.7266
22.491 | -103.7356
22.4945 | -103.7403
22.4981 | -103.7463
22.5017 | -103.7519
22.5052 | -103.7551
22.5088 | -103.7542
22.5124 | -103.7512
22.516 | -103.7427
22.5195 | -103.7334
22.5231 | -103.7258
22.5267 | -103.7177
22.5302 | -103.7074
22.5338 | -103.6945
22.5374 | -103.6819
22.5409 | -103.666
22.5445 | -103.6476
22.5481 | -103.629
22.5517 | -103.6098
22.5552 | -103.5902
22.5588 | -103.5662
22.5624 | -103.5376
22.5659 | -103.5088
22.5695 | -103.4815
22.5731 | -103.4551
22.5766 | -103.4286
22.5802 | -103.4032
22.5838 | -103.3739
22.5874 | -103.3413
22.5909 | -103.3094
22.5945 | -103.2805
22.5981 | -103.2534
22.6016 | -103.2254
22.6052 | -103.1972
22.6088 | -103.1702
22.6123 | -103.1413
22.6159 | -103.1119
22.6195 | -103.0849
22.6231 | -103.0601
22.6266 | -103.0367
22.6302 | -103.013
22.6338 | -102.9906
22.6373 | -102.968
22.6409 | -102.9448
22.6445 | -102.9236
22.648 | -102.9048
22.6516 | -102.888
22.6552 | -102.8713
22.6588 | -102.8624
22.6623 | -102.8471
22.6659 | -102.8325
22.6695 | -102.8201
22.673 | -102.8084
22.6766 | -102.7958
22.6802 | -102.785
22.6837 | -102.775
22.6873 | -102.7658
22.6909 | -102.7566
22.6945 | -102.7473
22.698 | -102.74
22.7016 | -102.7332
22.7052 | -102.7271
22.7087 | -102.7206
22.7123 | -102.7133
22.7159 | -102.7079
22.7194 | -102.7036
22.723 | -102.6975
22.7266 | -102.6907
22.7302 | -102.684
22.7337 | -102.6779
22.7373 | -102.6718
22.7409 | -102.665
22.7444 | -102.6586
22.748 | -102.6518
22.7516 | -102.6446
22.7551 | -102.6353
22.7587 | -102.6257
22.7623 | -102.6167
22.7659 | -102.6097
22.7694 | -102.604
22.773 | -102.5963
22.7766 | -102.5878
22.7801 | -102.5804
22.7837 | -102.5711
22.7873 | -102.5597
22.7908 | -102.5492
22.7944 | -102.5413
22.798 | -102.5316
22.8016 | -102.5228
22.8051 | -102.5131
22.8087 | -102.5008
22.8123 | -102.4892
22.8158 | -102.4799
22.8194 | -102.4698
22.823 | -102.4587
22.8265 | -102.4476
22.8301 | -102.4383
22.8337 | -102.4279
22.8373 | -102.4188
22.8408 | -102.409
22.8444 | -102.3965
22.848 | -102.3842
22.8515 | -102.373
22.8551 | -102.3622
22.8587 | -102.3513
22.8622 | -102.3406
22.8658 | -102.3299
22.8694 | -102.3211
22.873 | -102.3121
22.8765 | -102.3034
22.8801 | -102.2975
22.8837 | -102.2939
22.8872 | -102.2898
22.8908 | -102.2852
22.8944 | -102.2794
22.8979 | -102.2749
22.9015 | -102.2712
22.9051 | -102.2707
22.9087 | -102.2715
22.9122 | -102.2721
22.9158 | -102.2736
22.9194 | -102.2749
22.9229 | -102.2773
22.9265 | -102.2797
22.9301 | -102.2824
22.9336 | -102.288
22.9372 | -102.2952
22.9408 | -102.302
22.9444 | -102.3087
22.9479 | -102.3165
22.9515 | -102.3246
22.9551 | -102.3339
22.9586 | -102.3456
22.9622 | -102.3556
22.9658 | -102.3658
22.9693 | -102.3765
22.9729 | -102.388
22.9765 | -102.4007
22.9801 | -102.4129
22.9836 | -102.4262
22.9872 | -102.4399
22.9908 | -102.4525
22.9943 | -102.4655
22.9979 | -102.4797
23.0015 | -102.4953
23.005 | -102.5092
23.0086 | -102.5234
23.0122 | -102.5367
23.0158 | -102.5513
23.0193 | -102.564
23.0229 | -102.5728
23.0265 | -102.5848
23.03 | -102.5939
23.0336 | -102.605
23.0372 | -102.6135
23.0407 | -102.619
23.0443 | -102.6272
23.0479 | -102.6377
23.0515 | -102.6484
23.055 | -102.6577
23.0586 | -102.6631
23.0622 | -102.6678
23.0657 | -102.6744
23.0693 | -102.6798
23.0729 | -102.6843
23.0764 | -102.6881
23.08 | -102.6921
23.0836 | -102.6957
23.0872 | -102.6953
23.0907 | -102.6936
23.0943 | -102.6919
23.0979 | -102.6908
23.1014 | -102.6872
23.105 | -102.685
23.1086 | -102.684
23.1121 | -102.6815
23.1157 | -102.6792
23.1193 | -102.6744
23.1229 | -102.6674
23.1264 | -102.6599
23.13 | -102.6521
23.1336 | -102.6414
23.1371 | -102.6333
23.1407 | -102.6244
23.1443 | -102.6137
23.1478 | -102.6039
23.1514 | -102.594
23.155 | -102.5832
23.1586 | -102.5725
23.1621 | -102.5637
23.1657 | -102.5552
23.1693 | -102.5479
23.1728 | -102.5389
23.1764 | -102.5281
23.18 | -102.5189
23.1835 | -102.5105
23.1871 | -102.5028
23.1907 | -102.4968
23.1943 | -102.4919
23.1978 | -102.4838
23.2014 | -102.4749
23.205 | -102.4659
23.2085 | -102.4563
23.2121 | -102.446
23.2157 | -102.4363
23.2192 | -102.4254
23.2228 | -102.4171
23.2264 | -102.4087
23.23 | -102.399
23.2335 | -102.3888
23.2371 | -102.3798
23.2407 | -102.3701
23.2442 | -102.3596
23.2478 | -102.3484
23.2514 | -102.3378
23.2549 | -102.3299
23.2585 | -102.3233
23.2621 | -102.312
23.2657 | -102.3
23.2692 | -102.2908
23.2728 | -102.2821
23.2764 | -102.2726
23.2799 | -102.2634
23.2835 | -102.2542
23.2871 | -102.2421
23.2906 | -102.2313
23.2942 | -102.219
23.2978 | -102.2077
23.3014 | -102.1963
23.3049 | -102.1833
23.3085 | -102.1719
23.3121 | -102.1609
23.3156 | -102.1478
23.3192 | -102.1353
23.3228 | -102.1223
23.3263 | -102.1088
23.3299 | -102.0943
23.3335 | -102.0818
23.3371 | -102.0703
23.3406 | -102.0598
23.3442 | -102.0484
23.3478 | -102.0402
23.3513 | -102.0332
23.3549 | -102.0243
23.3585 | -102.0164
23.362 | -102.0079
23.3656 | -102.0002
23.3692 | -101.9946
23.3728 | -101.9876
23.3763 | -101.98
23.3799 | -101.9721
23.3835 | -101.9646
23.387 | -101.9585
23.3906 | -101.9532
23.3942 | -101.9468
23.3977 | -101.9422
23.4013 | -101.9395
23.4049 | -101.9366
23.4085 | -101.9347
23.412 | -101.9328
23.4156 | -101.9343
23.4192 | -101.9352
23.4227 | -101.9339
23.4263 | -101.9326
23.4299 | -101.9325
23.4334 | -101.93
23.437 | -101.9279
23.4406 | -101.9272
23.4442 | -101.927
23.4477 | -101.9253
23.4513 | -101.9265
23.4549 | -101.9262
23.4584 | -101.9264
23.462 | -101.93
23.4656 | -101.9324
23.4691 | -101.9325
23.4727 | -101.9304
23.4763 | -101.9281
23.4799 | -101.9274
23.4834 | -101.9258
23.487 | -101.9239
23.4906 | -101.9214
23.4941 | -101.9192
23.4977 | -101.9147
23.5013 | -101.9096
23.5048 | -101.9035
23.5084 | -101.8959
23.512 | -101.8875
23.5156 | -101.8793
23.5191 | -101.8709
23.5227 | -101.861
23.5263 | -101.8505
23.5298 | -101.841
23.5334 | -101.8311
23.537 | -101.8213
23.5405 | -101.8124
23.5441 | -101.8023
23.5477 | -101.7911
23.5513 | -101.7795
23.5548 | -101.7675
23.5584 | -101.755
23.562 | -101.7431
23.5655 | -101.7301
23.5691 | -101.7155
23.5727 | -101.7013
23.5762 | -101.6873
23.5798 | -101.6748
23.5834 | -101.6625
23.587 | -101.6485
23.5905 | -101.6348
23.5941 | -101.6248
23.5977 | -101.6153
23.6012 | -101.6043
23.6048 | -101.5945
23.6084 | -101.5863
23.6119 | -101.5785
23.6155 | -101.5719
23.6191 | -101.5692
23.6227 | -101.5674
23.6262 | -101.5656
23.6298 | -101.5622
23.6334 | -101.5628
23.6369 | -101.5657
23.6405 | -101.5692
23.6441 | -101.5736
23.6476 | -101.582
23.6512 | -101.5922
23.6548 | -101.6036
23.6584 | -101.616
23.6619 | -101.6295
23.6655 | -101.6442
23.6691 | -101.6604
23.6726 | -101.6775
23.6762 | -101.6967
23.6798 | -101.7148
23.6833 | -101.7349
23.6869 | -101.7561
23.6905 | -101.7885
23.6941 | -101.8137
23.6976 | -101.8402
23.7012 | -101.8683
23.7048 | -101.8944
23.7083 | -101.9218
23.7119 | -101.9517
23.7155 | -101.9845
23.719 | -102.017
23.7226 | -102.0495
23.7262 | -102.0807
23.7298 | -102.1121
23.7333 | -102.1429
23.7369 | -102.1743
23.7405 | -102.2082
23.744 | -102.2407
23.7476 | -102.2741
23.7512 | -102.3065
23.7547 | -102.3389
23.7583 | -102.3717
23.7619 | -102.4029
23.7655 | -102.4362
23.769 | -102.4693
23.7726 | -102.5019
23.7762 | -102.5341
23.7797 | -102.5624
23.7833 | -102.5918
23.7869 | -102.6196
23.7904 | -102.6453
23.794 | -102.6671
23.7976 | -102.6868
23.8012 | -102.7059
23.8047 | -102.7261
23.8083 | -102.7468
23.8119 | -102.7653
23.8154 | -102.7825
23.819 | -102.7994
23.8226 | -102.8125
23.8261 | -102.8249
23.8297 | -102.8368
23.8333 | -102.8493
23.8369 | -102.8599
23.8404 | -102.8702
23.844 | -102.8793
23.8476 | -102.885
23.8511 | -102.887
23.8547 | -102.8878
23.8583 | -102.8862
23.8618 | -102.8821
23.8654 | -102.8788
23.869 | -102.8749
23.8726 | -102.8685
23.8761 | -102.8602
23.8797 | -102.8518
23.8833 | -102.843
23.8868 | -102.8347
23.8904 | -102.8232
23.894 | -102.8126
23.8975 | -102.8007
23.9011 | -102.7882
23.9047 | -102.7756
23.9083 | -102.7618
23.9118 | -102.7475
23.9154 | -102.7331
23.919 | -102.7191
23.9225 | -102.7038
23.9261 | -102.6887
23.9297 | -102.6742
23.9332 | -102.6587
23.9368 | -102.6415
23.9404 | -102.6256
23.944 | -102.6103
23.9475 | -102.5951
23.9511 | -102.5803
23.9547 | -102.5648
23.9582 | -102.5505
23.9618 | -102.5365
23.9654 | -102.5245
23.9689 | -102.5096
23.9725 | -102.4955
23.9761 | -102.483
23.9797 | -102.4695
23.9832 | -102.455
23.9868 | -102.4393
23.9904 | -102.4231
23.9939 | -102.4107
23.9975 | -102.4009
24.0011 | -102.3902
24.0046 | -102.3776
24.0082 | -102.3655
24.0118 | -102.3556
24.0154 | -102.3484
24.0189 | -102.3408
24.0225 | -102.3343
24.0261 | -102.3274
24.0296 | -102.3217
24.0332 | -102.3152
24.0368 | -102.3085
24.0403 | -102.3046
24.0439 | -102.3017
24.0475 | -102.2981
24.0511 | -102.2947
24.0546 | -102.2916
24.0582 | -102.29
24.0618 | -102.2888
24.0653 | -102.2876
24.0689 | -102.2869
24.0725 | -102.2861
24.076 | -102.2858
24.0796 | -102.2853
24.0832 | -102.2861
24.0868 | -102.2852
24.0903 | -102.2848
24.0939 | -102.2851
24.0975 | -102.2866
24.101 | -102.2902
24.1046 | -102.2912
24.1082 | -102.2917
24.1117 | -102.2925
24.1153 | -102.2924
24.1189 | -102.2932
24.1225 | -102.2925
24.126 | -102.292
24.1296 | -102.2935
24.1332 | -102.2972
24.1367 | -102.3009
24.1403 | -102.3025
24.1439 | -102.3031
24.1474 | -102.305
24.151 | -102.3067
24.1546 | -102.3076
24.1582 | -102.3096
24.1617 | -102.3119
24.1653 | -102.3131
24.1689 | -102.3144
24.1724 | -102.3173
24.176 | -102.3215
24.1796 | -102.3249
24.1831 | -102.3307
24.1867 | -102.335
24.1903 | -102.3377
24.1939 | -102.3408
24.1974 | -102.3456
24.201 | -102.3508
24.2046 | -102.3555
24.2081 | -102.3608
24.2117 | -102.3661
24.2153 | -102.373
24.2188 | -102.3792
24.2224 | -102.3847
24.226 | -102.3895
24.2296 | -102.3929
24.2331 | -102.3967
24.2367 | -102.4004
24.2403 | -102.4045
24.2438 | -102.4076
24.2474 | -102.4122
24.251 | -102.418
24.2545 | -102.4213
24.2581 | -102.4228
24.2617 | -102.4219
24.2653 | -102.4208
24.2688 | -102.4143
24.2724 | -102.4024
24.276 | -102.3873
24.2795 | -102.3667
24.2831 | -102.3446
24.2867 | -102.327
24.2902 | -102.3136
24.2938 | -102.3008
24.2974 | -102.2863
24.301 | -102.2698
24.3045 | -102.2534
24.3081 | -102.2366
24.3117 | -102.2174
24.3152 | -102.1959
24.3188 | -102.1741
24.3224 | -102.1516
24.3259 | -102.1283
24.3295 | -102.1059
24.3331 | -102.0826
24.3366 | -102.056
24.3402 | -102.0268
24.3438 | -101.9976
24.3474 | -101.9701
24.3509 | -101.9411
24.3545 | -101.91
24.3581 | -101.8786
24.3616 | -101.8483
24.3652 | -101.8181
24.3688 | -101.789
24.3724 | -101.7595
24.3759 | -101.7316
24.3795 | -101.7048
24.3831 | -101.6778
24.3866 | -101.6521
24.3902 | -101.6278
24.3938 | -101.6054
24.3973 | -101.5848
24.4009 | -101.5671
24.4045 | -101.5516
24.408 | -101.5384
24.4116 | -101.5265
24.4152 | -101.5166
24.4188 | -101.5086
24.4223 | -101.5018
24.4259 | -101.4974
24.4295 | -101.4971
24.433 | -101.5001
24.4366 | -101.5066
24.4402 | -101.5155
24.4438 | -101.5255
24.4473 | -101.539
24.4509 | -101.5566
24.4545 | -101.5763
24.458 | -101.5985
24.4616 | -101.6245
24.4652 | -101.655
24.4687 | -101.6862
24.4723 | -101.7186
24.4759 | -101.7536
24.4794 | -101.7896
24.483 | -101.8279
24.4866 | -101.8693
24.4902 | -101.9135
24.4937 | -101.9602
24.4973 | -102.0081
24.5009 | -102.0561
24.5044 | -102.1068
24.508 | -102.1596
24.5116 | -102.2127
24.5151 | -102.2658
24.5187 | -102.3198
24.5223 | -102.3738
24.5259 | -102.4267
24.5294 | -102.4809
24.533 | -102.5371
24.5366 | -102.5918
24.5401 | -102.6454
24.5437 | -102.6974
24.5473 | -102.748
24.5509 | -102.798
24.5544 | -102.8459
24.558 | -102.8916
24.5616 | -102.9348
24.5651 | -102.9764
24.5687 | -103.0153
24.5723 | -103.0526
24.5758 | -103.0895
24.5794 | -103.1231
24.583 | -103.1495
24.5865 | -103.1633
24.5901 | -103.1566
24.5937 | -103.1361
24.5973 | -103.1123
24.6008 | -103.0924
24.6044 | -103.0773
24.608 | -103.0634
24.6115 | -103.0434
24.6151 | -103.0226
24.6187 | -103.0043
24.6223 | -102.9841
24.6258 | -102.963
24.6294 | -102.9335
24.633 | -102.8909
24.6365 | -102.8391
24.6401 | -102.7827
24.6437 | -102.7237
24.6472 | -102.6631
24.6508 | -102.5995
24.6544 | -102.5319
24.6579 | -102.4643
24.6615 | -102.3954
24.6651 | -102.3238
24.6687 | -102.2536
24.6722 | -102.182
24.6758 | -102.1082
24.6794 | -102.0339
24.6829 | -101.9586
24.6865 | -101.8842
24.6901 | -101.8103
24.6937 | -101.737
24.6972 | -101.6655
24.7008 | -101.5962
24.7044 | -101.5278
24.7079 | -101.4607
24.7115 | -101.3958
24.7151 | -101.3325
24.7186 | -101.2931
24.7222 | -101.2389
24.7258 | -101.1876
24.7294 | -101.1402
24.7329 | -101.0972
24.7365 | -101.0582
24.7401 | -101.0224
24.7436 | -100.99
24.7472 | -100.9617
24.7508 | -100.9372
24.7543 | -100.9174
24.7579 | -100.9021
24.7615 | -100.8911
24.765 | -100.8851
24.7686 | -100.8835
24.7722 | -100.8871
24.7758 | -100.8947
24.7793 | -100.9065
24.7829 | -100.9221
24.7865 | -100.9418
24.79 | -100.9662
24.7936 | -100.9957
24.7972 | -101.0286
24.8008 | -101.0649
24.8043 | -101.1052
24.8079 | -101.1491
24.8115 | -101.1954
24.815 | -101.245
24.8186 | -101.2987
24.8222 | -101.3558
24.8257 | -101.4154
24.8293 | -101.4779
24.8329 | -101.5423
24.8364 | -101.608
24.84 | -101.6752
24.8436 | -101.7436
24.8472 | -101.8128
24.8507 | -101.8845
24.8543 | -101.9571
24.8579 | -102.0286
24.8614 | -102.0996
24.865 | -102.1712
24.8686 | -102.2432
24.8722 | -102.3156
24.8757 | -102.3859
24.8793 | -102.4556
24.8829 | -102.5246
24.8864 | -102.5924
24.89 | -102.6587
24.8936 | -102.7243
24.8971 | -102.787
24.9007 | -102.8463
24.9043 | -102.9037
24.9078 | -102.9599
24.9114 | -103.0146
24.915 | -103.0665
24.9186 | -103.1159
24.9221 | -103.1618
24.9257 | -103.2053
24.9293 | -103.2482
24.9328 | -103.2921
24.9364 | -103.3427
24.94 | -103.4059
24.9435 | -103.4702
24.9471 | -103.5244
24.9507 | -103.5639
24.9543 | -103.5895
24.9578 | -103.6089
24.9614 | -103.6295
24.965 | -103.6477
24.9685 | -103.6591
24.9721 | -103.6693
24.9757 | -103.6793
24.9792 | -103.6877
24.9828 | -103.6966
24.9864 | -103.7059
24.99 | -103.7137
24.9935 | -103.721
24.9971 | -103.7284
25.0007 | -103.7359
25.0042 | -103.7421
25.0078 | -103.7477
25.0114 | -103.7558
25.0149 | -103.7639
25.0185 | -103.7692
25.0221 | -103.7753
25.0257 | -103.7823
25.0292 | -103.7896
25.0328 | -103.7971
25.0364 | -103.8044
25.0399 | -103.8129
25.0435 | -103.819
25.0471 | -103.8236
25.0506 | -103.8264
25.0542 | -103.8309
25.0578 | -103.8359
25.0614 | -103.8408
25.0649 | -103.8454
25.0685 | -103.8499
25.0721 | -103.8522
25.0756 | -103.8555
25.0792 | -103.8597
25.0828 | -103.8626
25.0863 | -103.864
25.0899 | -103.8651
25.0935 | -103.8669
25.0971 | -103.8695
25.1006 | -103.8724
25.1042 | -103.8746
25.1078 | -103.876
25.1113 | -103.8746
25.1149 | -103.8755
25.1185 | -103.8767
25.122 | -103.8775
25.1256 | -103.8764
25.1292 | -103.8776
25.1328 | -103.8791
25.1363 | -103.8813
25.1399 | -103.8839
25.1435 | -103.8852
25.147 | -103.8858
25.1506 | -103.8873
25.1542 | -103.8863
25.1577 | -103.8852
25.1613 | -103.8867
25.1649 | -103.8885
25.1685 | -103.8886
25.172 | -103.8878
25.1756 | -103.887
25.1792 | -103.8864
25.1827 | -103.887
25.1863 | -103.8879
25.1899 | -103.8882
25.1934 | -103.8869
25.197 | -103.8842
25.2006 | -103.883
25.2042 | -103.8824
25.2077 | -103.8808
25.2113 | -103.8792
25.2149 | -103.8779
25.2184 | -103.8778
25.222 | -103.8773
25.2256 | -103.8768
25.2291 | -103.8756
25.2327 | -103.8751
25.2363 | -103.8733
25.2399 | -103.8716
25.2434 | -103.8693
25.247 | -103.8675
25.2506 | -103.8657
25.2541 | -103.8639
25.2577 | -103.8628
25.2613 | -103.8603
25.2648 | -103.8562
25.2684 | -103.8523
25.272 | -103.852
25.2756 | -103.8533
25.2791 | -103.8542
25.2827 | -103.8528
25.2863 | -103.8496
25.2898 | -103.8474
25.2934 | -103.845
25.297 | -103.8417
25.3005 | -103.8404
25.3041 | -103.841
25.3077 | -103.8436
25.3113 | -103.8457
25.3148 | -103.8468
25.3184 | -103.8457
25.322 | -103.8455
25.3255 | -103.8463
25.3291 | -103.8458
25.3327 | -103.8467
25.3362 | -103.8479
25.3398 | -103.848
25.3434 | -103.8472
25.347 | -103.8483
25.3505 | -103.8514
25.3541 | -103.8533
25.3577 | -103.8558
25.3612 | -103.8606
25.3648 | -103.8636
25.3684 | -103.8646
25.3719 | -103.8659
25.3755 | -103.8659
25.3791 | -103.8665
25.3827 | -103.868
25.3862 | -103.8707
25.3898 | -103.8725
25.3934 | -103.8744
25.3969 | -103.8754
25.4005 | -103.8777
25.4041 | -103.8807
25.4076 | -103.8835
25.4112 | -103.8843
25.4148 | -103.8845
25.4184 | -103.885
25.4219 | -103.8848
25.4255 | -103.8831
25.4291 | -103.8821
25.4326 | -103.8814
25.4362 | -103.8793
25.4398 | -103.8773
25.4433 | -103.8763
25.4469 | -103.8751
25.4505 | -103.8717
25.4541 | -103.8668
25.4576 | -103.8621
25.4612 | -103.8583
25.4648 | -103.8534
25.4683 | -103.8503
25.4719 | -103.8464
25.4755 | -103.8429
25.479 | -103.8391
25.4826 | -103.8346
25.4862 | -103.8289
25.4898 | -103.8221
25.4933 | -103.8137
25.4969 | -103.8047
25.5005 | -103.7994
25.504 | -103.7948
25.5076 | -103.7892
25.5112 | -103.7847
25.5147 | -103.7799
25.5183 | -103.7723
25.5219 | -103.7661
25.5255 | -103.7604
25.529 | -103.7557
25.5326 | -103.7525
25.5362 | -103.748
25.5397 | -103.7426
25.5433 | -103.7382
25.5469 | -103.7346
25.5504 | -103.7296
25.554 | -103.7256
25.5576 | -103.7229
25.5612 | -103.7191
25.5647 | -103.716
25.5683 | -103.7138
25.5719 | -103.7127
25.5754 | -103.7095
25.579 | -103.708
25.5826 | -103.7063
25.5861 | -103.7059
25.5897 | -103.7074
25.5933 | -103.7099
25.5969 | -103.7112
25.6004 | -103.7134
25.604 | -103.7166
25.6076 | -103.7199
25.6111 | -103.7225
25.6147 | -103.7258
25.6183 | -103.7305
25.6218 | -103.7363
25.6254 | -103.744
25.629 | -103.7522
25.6326 | -103.7581
25.6361 | -103.7638
25.6397 | -103.7692
25.6433 | -103.7755
25.6468 | -103.7829
25.6504 | -103.7908
25.654 | -103.7974
25.6575 | -103.8052
25.6611 | -103.812
25.6647 | -103.8171
25.6683 | -103.823
25.6718 | -103.8293
25.6754 | -103.8366
25.679 | -103.8466
25.6825 | -103.8553
25.6861 | -103.8629
25.6897 | -103.87
25.6932 | -103.8761
25.6968 | -103.8817
25.7004 | -103.8877
25.704 | -103.8955
25.7075 | -103.9011
25.7111 | -103.9062
25.7147 | -103.9132
25.7182 | -103.9209
25.7218 | -103.9276
25.7254 | -103.9342
25.7289 | -103.9404
25.7325 | -103.946
25.7361 | -103.9538
25.7397 | -103.9617
25.7432 | -103.9685
25.7468 | -103.9746
25.7504 | -103.9804
25.7539 | -103.987
25.7575 | -103.9933
25.7611 | -104.0006
25.7646 | -104.0082
25.7682 | -104.015
25.7718 | -104.0207
25.7754 | -104.0281
25.7789 | -104.0366
25.7825 | -104.0447
25.7861 | -104.0524
25.7896 | -104.0604
25.7932 | -104.0681
25.7968 | -104.0756
25.8003 | -104.0822
25.8039 | -104.0897
25.8075 | -104.0968
25.8111 | -104.1043
25.8146 | -104.1121
25.8182 | -104.121
25.8218 | -104.1295
25.8253 | -104.1376
25.8289 | -104.1455
25.8325 | -104.1538
25.836 | -104.1627
25.8396 | -104.1698
25.8432 | -104.1757
25.8468 | -104.1847
25.8503 | -104.1939
25.8539 | -104.2037
25.8575 | -104.2139
25.861 | -104.2232
25.8646 | -104.2322
25.8682 | -104.2433
25.8717 | -104.2549
25.8753 | -104.2664
25.8789 | -104.2781
25.8825 | -104.2903
25.886 | -104.3014
25.8896 | -104.3122
25.8932 | -104.3224
25.8967 | -104.3327
25.9003 | -104.343
25.9039 | -104.3543
25.9074 | -104.3653
25.911 | -104.3772
25.9146 | -104.3917
25.9182 | -104.4052
25.9217 | -104.4161
25.9253 | -104.4277
25.9289 | -104.4394
25.9324 | -104.4512
25.936 | -104.463
25.9396 | -104.4745
25.9431 | -104.4839
25.9467 | -104.4947
25.9503 | -104.5053
25.9539 | -104.5142
25.9574 | -104.5238
25.961 | -104.5332
25.9646 | -104.5422
25.9681 | -104.5514
25.9717 | -104.5602
25.9753 | -104.5682
25.9788 | -104.575
25.9824 | -104.5837
25.986 | -104.5928
25.9896 | -104.5993
25.9931 | -104.6048
25.9967 | -104.611
26.0003 | -104.6172
26.0038 | -104.6247
26.0074 | -104.6326
26.011 | -104.6386
26.0145 | -104.6423
26.0181 | -104.6465
26.0217 | -104.6503
26.0253 | -104.6533
26.0288 | -104.6576
26.0324 | -104.6601
26.036 | -104.6619
26.0395 | -104.6632
26.0431 | -104.6645
26.0467 | -104.6675
26.0502 | -104.6705
26.0538 | -104.6729
26.0574 | -104.6747
26.061 | -104.6767
26.0645 | -104.6773
26.0681 | -104.6767
26.0717 | -104.6772
26.0752 | -104.6775
26.0788 | -104.6774
26.0824 | -104.6747
26.0859 | -104.671
26.0895 | -104.6676
26.0931 | -104.6655
26.0967 | -104.6636
26.1002 | -104.6619
26.1038 | -104.6594
26.1074 | -104.6562
26.1109 | -104.6536
26.1145 | -104.6518
26.1181 | -104.6495
26.1216 | -104.6468
26.1252 | -104.6439
26.1288 | -104.6392
26.1324 | -104.6334
26.1359 | -104.627
26.1395 | -104.6222
26.1431 | -104.6175
26.1466 | -104.6118
26.1502 | -104.6062
26.1538 | -104.6011
26.1573 | -104.5969
26.1609 | -104.5927
26.1645 | -104.587
26.1681 | -104.5837
26.1716 | -104.5798
26.1752 | -104.5757
26.1788 | -104.573
26.1823 | -104.5702
26.1859 | -104.5651
26.1895 | -104.5597
26.193 | -104.5554
26.1966 | -104.5536
26.2002 | -104.5509
26.2038 | -104.5465
26.2073 | -104.5431
26.2109 | -104.5415
26.2145 | -104.538
26.218 | -104.5349
26.2216 | -104.53
26.2252 | -104.5263
26.2287 | -104.524
26.2323 | -104.5195
26.2359 | -104.5141
26.2395 | -104.5092
26.243 | -104.5062
26.2466 | -104.5036
26.2502 | -104.5014
26.2537 | -104.4986
26.2573 | -104.4959
26.2609 | -104.4942
26.2644 | -104.4931
26.268 | -104.4915
26.2716 | -104.4904
26.2752 | -104.4887
26.2787 | -104.4858
26.2823 | -104.4833
26.2859 | -104.482
26.2894 | -104.4795
26.293 | -104.4775
26.2966 | -104.4752
26.3001 | -104.4728
26.3037 | -104.4716
26.3073 | -104.47
26.3109 | -104.4691
26.3144 | -104.4687
26.318 | -104.4672
26.3216 | -104.4659
26.3251 | -104.4645
26.3287 | -104.4619
26.3323 | -104.4625
26.3358 | -104.4631
26.3394 | -104.4621
26.343 | -104.4615
26.3466 | -104.463
26.3501 | -104.4648
26.3537 | -104.4623
26.3573 | -104.4608
26.3608 | -104.4604
26.3644 | -104.4583
26.368 | -104.4556
26.3715 | -104.4551
26.3751 | -104.4544
26.3787 | -104.4533
26.3823 | -104.4539
26.3858 | -104.4558
26.3894 | -104.458
26.393 | -104.459
26.3965 | -104.459
26.4001 | -104.4598
26.4037 | -104.4614
26.4072 | -104.4624
26.4108 | -104.4633
26.4144 | -104.4635
26.418 | -104.4633
26.4215 | -104.4631
26.4251 | -104.4628
26.4287 | -104.4645
26.4322 | -104.4658
26.4358 | -104.4671
26.4394 | -104.4666
26.4429 | -104.4659
26.4465 | -104.4657
26.4501 | -104.4655
26.4537 | -104.4664
26.4572 | -104.468
26.4608 | -104.4695
26.4644 | -104.4696
26.4679 | -104.47
26.4715 | -104.4686
26.4751 | -104.4676
26.4786 | -104.468
26.4822 | -104.4675
26.4858 | -104.4659
26.4894 | -104.4655
26.4929 | -104.4661
26.4965 | -104.4651
26.5001 | -104.463
26.5036 | -104.463
26.5072 | -104.4632
26.5108 | -104.4627
26.5143 | -104.4619
26.5179 | -104.4618
26.5215 | -104.4622
26.5251 | -104.4643
26.5286 | -104.4647
26.5322 | -104.4651
26.5358 | -104.4671
26.5393 | -104.4676
26.5429 | -104.4674
26.5465 | -104.4678
26.55 | -104.4675
26.5536 | -104.4688
26.5572 | -104.4711
26.5608 | -104.4731
26.5643 | -104.4732
26.5679 | -104.4731
26.5715 | -104.4734
26.575 | -104.4737
26.5786 | -104.4733
26.5822 | -104.4719
26.5857 | -104.4719
26.5893 | -104.471
26.5929 | -104.4721
26.5965 | -104.4721
26.6 | -104.4718
26.6036 | -104.4717
26.6072 | -104.4716
26.6107 | -104.4724
26.6143 | -104.4719
26.6179 | -104.4704
26.6214 | -104.47
26.625 | -104.4706
26.6286 | -104.4698
26.6322 | -104.4685
26.6357 | -104.4672
26.6393 | -104.4672
26.6429 | -104.4668
26.6464 | -104.467
26.65 | -104.4684
26.6536 | -104.4711
26.6571 | -104.4737
26.6607 | -104.4757
26.6643 | -104.4786
26.6679 | -104.481
26.6714 | -104.4839
26.675 | -104.4876
26.6786 | -104.4916
26.6821 | -104.4956
26.6857 | -104.4995
26.6893 | -104.5036
26.6928 | -104.5093
26.6964 | -104.5159
26.7 | -104.5212
26.7036 | -104.5279
26.7071 | -104.5337
26.7107 | -104.5394
26.7143 | -104.5452
26.7178 | -104.5521
26.7214 | -104.5603
26.725 | -104.5696
26.7285 | -104.5794
26.7321 | -104.5879
26.7357 | -104.5955
26.7393 | -104.6037
26.7428 | -104.6129
26.7464 | -104.6197
26.75 | -104.6261
26.7535 | -104.6336
26.7571 | -104.6403
26.7607 | -104.6464
26.7642 | -104.6534
26.7678 | -104.6604
26.7714 | -104.6679
26.775 | -104.6761
26.7785 | -104.6821
26.7821 | -104.6887
26.7857 | -104.6949
26.7892 | -104.7015
26.7928 | -104.7077
26.7964 | -104.7131
26.7999 | -104.7177
26.8035 | -104.7233
26.8071 | -104.7274
26.8107 | -104.7303
26.8142 | -104.7351
26.8178 | -104.7413
26.8214 | -104.7468
26.8249 | -104.7506
26.8285 | -104.7531
26.8321 | -104.7559
26.8356 | -104.7592
26.8392 | -104.7633
26.8428 | -104.7672
26.8464 | -104.7707
26.8499 | -104.7749
26.8535 | -104.7759
26.8571 | -104.7752
26.8606 | -104.7756
26.8642 | -104.7756
26.8678 | -104.7745
26.8713 | -104.7738
26.8749 | -104.7736
26.8785 | -104.7727
26.8821 | -104.7712
26.8856 | -104.7705
26.8892 | -104.7686
26.8928 | -104.766
26.8963 | -104.7626
26.8999 | -104.7587
26.9035 | -104.7541
26.907 | -104.7502
26.9106 | -104.746
26.9142 | -104.7419
26.9178 | -104.7371
26.9213 | -104.7341
26.9249 | -104.7311
26.9285 | -104.7265
26.932 | -104.7224
26.9356 | -104.7191
26.9392 | -104.7161
26.9427 | -104.7132
26.9463 | -104.7099
26.9499 | -104.7075
26.9535 | -104.7059
26.957 | -104.7047
26.9606 | -104.7035
26.9642 | -104.7021
26.9677 | -104.7014
26.9713 | -104.7002
26.9749 | -104.6983
26.9784 | -104.6961
26.982 | -104.6938
26.9856 | -104.6927
26.9892 | -104.6923
26.9927 | -104.6931
26.9963 | -104.6952
26.9999 | -104.6967
27.0034 | -104.6979
27.007 | -104.7015
27.0106 | -104.7028
27.0141 | -104.7052
27.0177 | -104.7073
27.0213 | -104.7082
27.0249 | -104.7091
27.0284 | -104.7098
27.032 | -104.7124
27.0356 | -104.7156
27.0391 | -104.7187
27.0427 | -104.7238
27.0463 | -104.7293
27.0498 | -104.7334
27.0534 | -104.7363
27.057 | -104.7387
27.0606 | -104.7421
27.0641 | -104.7462
27.0677 | -104.7498
27.0713 | -104.753
27.0748 | -104.7555
27.0784 | -104.7575
27.082 | -104.7585
27.0855 | -104.7589
27.0891 | -104.7599
27.0927 | -104.761
27.0963 | -104.7616
27.0998 | -104.7616
27.1034 | -104.7619
27.107 | -104.7631
27.1105 | -104.7661
27.1141 | -104.768
27.1177 | -104.7688
27.1212 | -104.7688
27.1248 | -104.7687
27.1284 | -104.7675
27.132 | -104.7649
27.1355 | -104.7597
27.1391 | -104.7556
27.1427 | -104.7519
27.1462 | -104.7489
27.1498 | -104.7434
27.1534 | -104.7349
27.1569 | -104.7265
27.1605 | -104.7169
27.1641 | -104.7099
27.1677 | -104.7038
27.1712 | -104.6957
27.1748 | -104.6876
27.1784 | -104.6786
27.1819 | -104.669
27.1855 | -104.6589
27.1891 | -104.6483
27.1926 | -104.6383
27.1962 | -104.6261
27.1998 | -104.6123
27.2034 | -104.5982
27.2069 | -104.5851
27.2105 | -104.5733
27.2141 | -104.5615
27.2176 | -104.548
27.2212 | -104.5335
27.2248 | -104.5209
27.2283 | -104.5089
27.2319 | -104.4959
27.2355 | -104.4845
27.2391 | -104.4732
27.2426 | -104.4607
27.2462 | -104.4496
27.2498 | -104.4391
27.2533 | -104.4285
27.2569 | -104.4184
27.2605 | -104.4097
27.264 | -104.4004
27.2676 | -104.3912
27.2712 | -104.3826
27.2748 | -104.3734
27.2783 | -104.365
27.2819 | -104.3563
27.2855 | -104.3474
27.289 | -104.339
27.2926 | -104.3309
27.2962 | -104.3284
27.2997 | -104.3208
27.3033 | -104.3154
27.3069 | -104.3102
27.3105 | -104.3052
27.314 | -104.3008
27.3176 | -104.2959
27.3212 | -104.2922
27.3247 | -104.2879
27.3283 | -104.284
27.3319 | -104.2785
27.3354 | -104.2741
27.339 | -104.2708
27.3426 | -104.2663
27.3462 | -104.2606
27.3497 | -104.2553
27.3533 | -104.2488
27.3569 | -104.2429
27.3604 | -104.2368
27.364 | -104.2308
27.3676 | -104.2244
27.3711 | -104.2182
27.3747 | -104.212
27.3783 | -104.2048
27.3819 | -104.1976
27.3854 | -104.19
27.389 | -104.1831
27.3926 | -104.1755
27.3961 | -104.1684
27.3997 | -104.1606
27.4033 | -104.1522
27.4068 | -104.1457
27.4104 | -104.1379
27.414 | -104.1299
27.4176 | -104.1217
27.4211 | -104.1135
27.4247 | -104.1041
27.4283 | -104.0929
27.4318 | -104.0832
27.4354 | -104.0743
27.439 | -104.0651
27.4425 | -104.0565
27.4461 | -104.048
27.4497 | -104.0383
27.4533 | -104.0289
27.4568 | -104.0204
27.4604 | -104.0105
27.464 | -103.9999
27.4675 | -103.9908
27.4711 | -103.9833
27.4747 | -103.9749
27.4782 | -103.9664
27.4818 | -103.9562
27.4854 | -103.9441
27.489 | -103.9322
27.4925 | -103.9189
27.4961 | -103.9044
27.4997 | -103.8926
27.5032 | -103.8827
27.5068 | -103.8724
27.5104 | -103.8605
27.5139 | -103.8492
27.5175 | -103.8377
27.5211 | -103.8277
27.5247 | -103.816
27.5282 | -103.8024
27.5318 | -103.7886
27.5354 | -103.7778
27.5389 | -103.7676
27.5425 | -103.7567
27.5461 | -103.7458
27.5496 | -103.7323
27.5532 | -103.7191
27.5568 | -103.706
27.5604 | -103.6921
27.5639 | -103.6765
27.5675 | -103.6634
27.5711 | -103.651
27.5746 | -103.6365
27.5782 | -103.6223
27.5818 | -103.6094
27.5853 | -103.5973
27.5889 | -103.5834
27.5925 | -103.5704
27.5961 | -103.5558
27.5996 | -103.5427
27.6032 | -103.5303
27.6068 | -103.5181
27.6103 | -103.505
27.6139 | -103.492
27.6175 | -103.4792
27.621 | -103.4672
27.6246 | -103.4548
27.6282 | -103.4435
27.6318 | -103.4318
27.6353 | -103.4198
27.6389 | -103.4097
27.6425 | -103.3997
27.646 | -103.3895
27.6496 | -103.38
27.6532 | -103.3721
27.6567 | -103.3651
27.6603 | -103.3581
27.6639 | -103.351
27.6675 | -103.3452
27.671 | -103.3407
27.6746 | -103.3358
27.6782 | -103.3321
27.6817 | -103.3274
27.6853 | -103.3232
27.6889 | -103.3203
27.6924 | -103.317
27.696 | -103.3155
27.6996 | -103.315
27.7032 | -103.3151
27.7067 | -103.3152
27.7103 | -103.3152
27.7139 | -103.3167
27.7174 | -103.3201
27.721 | -103.3247
27.7246 | -103.3279
27.7281 | -103.3314
27.7317 | -103.3365
27.7353 | -103.3431
27.7389 | -103.3484
27.7424 | -103.3535
27.746 | -103.3577
27.7496 | -103.3637
27.7531 | -103.37
27.7567 | -103.377
27.7603 | -103.3823
27.7638 | -103.3859
27.7674 | -103.3908
27.771 | -103.3963
27.7746 | -103.402
27.7781 | -103.4082
27.7817 | -103.4128
27.7853 | -103.4165
27.7888 | -103.4206
27.7924 | -103.4258
27.796 | -103.4297
27.7995 | -103.4335
27.8031 | -103.4365
27.8067 | -103.4376
27.8103 | -103.4314
27.8138 | -103.4298
27.8174 | -103.4295
27.821 | -103.4304
27.8245 | -103.4287
27.8281 | -103.4255
27.8317 | -103.4211
27.8352 | -103.416
27.8388 | -103.4108
27.8424 | -103.4063
27.846 | -103.4014
27.8495 | -103.3944
27.8531 | -103.3877
27.8567 | -103.3804
27.8602 | -103.3715
27.8638 | -103.3612
27.8674 | -103.3508
27.8709 | -103.3397
27.8745 | -103.3298
27.8781 | -103.3212
27.8817 | -103.3113
27.8852 | -103.302
27.8888 | -103.2932
27.8924 | -103.2837
27.8959 | -103.2734
27.8995 | -103.2621
27.9031 | -103.2514
27.9066 | -103.2398
27.9102 | -103.2283
27.9138 | -103.2181
27.9174 | -103.2085
27.9209 | -103.2004
27.9245 | -103.1922
27.9281 | -103.1837
27.9316 | -103.1756
27.9352 | -103.1676
27.9388 | -103.1612
27.9423 | -103.1553
27.9459 | -103.1497
27.9495 | -103.1448
27.9531 | -103.1418
27.9566 | -103.1407
27.9602 | -103.1375
27.9638 | -103.1348
27.9673 | -103.1358
27.9709 | -103.1376
27.9745 | -103.1396
27.978 | -103.1413
27.9816 | -103.1442
27.9852 | -103.1477
27.9888 | -103.1522
27.9923 | -103.1566
27.9959 | -103.1626
27.9995 | -103.1692
28.003 | -103.1766