## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 215 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 213
Hood | 02-92B/P | 152
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1011.9 | hPa |  | PRC PRO 01
Humidity | 41.9 | %/Hr
Internal temperature of calibration bench | 23.9 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1007 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 25.1 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.02 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.04 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3144 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 17/04/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 19/04/2018
Engineer | Anthony HUE | Approval |  | 23/04/2018
CEO | Robert GLEMOT | Approval |  | 23/04/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3759 | 57.5377
0.012209 | -35.9546 | 38.3627
0.018313 | -34.8965 | 28.2222
0.024417 | -34.4322 | 21.9721
0.030521 | -34.3198 | 18.2124
0.036626 | -34.1743 | 15.1665
0.04273 | -34.0785 | 12.8043
0.048834 | -34.0321 | 11.2803
0.054938 | -33.9663 | 10.1475
0.061043 | -33.9051 | 9.0707
0.067147 | -33.8946 | 8.2566
0.073251 | -33.895 | 7.5029
0.079355 | -33.9126 | 6.8433
0.08546 | -33.9258 | 6.2497
0.091564 | -33.9205 | 5.8094
0.097668 | -33.9024 | 5.5588
0.10377 | -33.8805 | 5.3109
0.10988 | -33.8632 | 5.0347
0.11598 | -33.8536 | 4.744
0.12209 | -33.8548 | 4.4624
0.12819 | -33.8642 | 4.1811
0.13429 | -33.8747 | 3.8958
0.1404 | -33.878 | 3.621
0.1465 | -33.88 | 3.3663
0.15261 | -33.8774 | 3.1483
0.15871 | -33.8712 | 2.9868
0.16482 | -33.8631 | 2.883
0.17092 | -33.8605 | 2.8409
0.17702 | -33.858 | 2.7664
0.18313 | -33.8568 | 2.686
0.18923 | -33.8561 | 2.5753
0.19534 | -33.8545 | 2.4474
0.20144 | -33.853 | 2.3048
0.20754 | -33.8499 | 2.1841
0.21365 | -33.8486 | 2.0695
0.21975 | -33.8479 | 1.9716
0.22586 | -33.8488 | 1.8843
0.23196 | -33.851 | 1.812
0.23807 | -33.8548 | 1.7492
0.24417 | -33.8586 | 1.6909
0.25027 | -33.8623 | 1.6288
0.25638 | -33.8645 | 1.5672
0.26248 | -33.8662 | 1.5008
0.26859 | -33.867 | 1.4356
0.27469 | -33.8672 | 1.3648
0.2808 | -33.868 | 1.2912
0.2869 | -33.8686 | 1.2203
0.293 | -33.8694 | 1.153
0.29911 | -33.8701 | 1.0927
0.30521 | -33.8708 | 1.0334
0.31132 | -33.8714 | 0.99395
0.31742 | -33.8724 | 0.9505
0.32353 | -33.8725 | 0.91488
0.32963 | -33.8732 | 0.88525
0.33573 | -33.874 | 0.86444
0.34184 | -33.8752 | 0.84267
0.34794 | -33.8771 | 0.81919
0.35405 | -33.8791 | 0.79514
0.36015 | -33.8814 | 0.76987
0.36626 | -33.8832 | 0.7425
0.37236 | -33.8852 | 0.70609
0.37846 | -33.8872 | 0.66216
0.38457 | -33.8881 | 0.60512
0.39067 | -33.889 | 0.55852
0.39678 | -33.8896 | 0.50671
0.40288 | -33.8899 | 0.45267
0.40899 | -33.8898 | 0.40073
0.41509 | -33.889 | 0.34969
0.42119 | -33.8889 | 0.30626
0.4273 | -33.8887 | 0.25987
0.4334 | -33.8885 | 0.21367
0.43951 | -33.8887 | 0.17004
0.44561 | -33.8894 | 0.13014
0.45172 | -33.8909 | 0.098033
0.45782 | -33.8916 | 0.060628
0.46392 | -33.8927 | 0.024129
0.47003 | -33.8937 | -0.0086793
0.47613 | -33.8944 | -0.043192
0.48224 | -33.8949 | -0.078455
0.48834 | -33.8939 | -0.11373
0.49445 | -33.8937 | -0.14942
0.50055 | -33.8931 | -0.18193
0.50665 | -33.8923 | -0.2166
0.51276 | -33.8912 | -0.25106
0.51886 | -33.8902 | -0.28539
0.52497 | -33.8889 | -0.31007
0.53107 | -33.888 | -0.33987
0.53717 | -33.8872 | -0.36849
0.54328 | -33.8863 | -0.39182
0.54938 | -33.8861 | -0.41474
0.55549 | -33.886 | -0.4361
0.56159 | -33.8869 | -0.44818
0.5677 | -33.8872 | -0.4633
0.5738 | -33.8882 | -0.48013
0.5799 | -33.8889 | -0.49558
0.58601 | -33.8899 | -0.5063
0.59211 | -33.8915 | -0.5133
0.59822 | -33.8928 | -0.52339
0.60432 | -33.8942 | -0.53597
0.61043 | -33.8955 | -0.54465
0.61653 | -33.8967 | -0.55562
0.62263 | -33.8982 | -0.56563
0.62874 | -33.8988 | -0.57699
0.63484 | -33.8996 | -0.58695
0.64095 | -33.9001 | -0.5978
0.64705 | -33.9005 | -0.60828
0.65316 | -33.9006 | -0.62129
0.65926 | -33.9007 | -0.63755
0.66536 | -33.8996 | -0.65525
0.67147 | -33.8992 | -0.66895
0.67757 | -33.8988 | -0.68371
0.68368 | -33.898 | -0.70288
0.68978 | -33.8972 | -0.72033
0.69589 | -33.8963 | -0.73955
0.70199 | -33.8953 | -0.76232
0.70809 | -33.8944 | -0.78167
0.7142 | -33.8936 | -0.8023
0.7203 | -33.8925 | -0.82168
0.72641 | -33.8918 | -0.84429
0.73251 | -33.8911 | -0.86441
0.73862 | -33.8913 | -0.88627
0.74472 | -33.8909 | -0.90648
0.75082 | -33.8906 | -0.92873
0.75693 | -33.8909 | -0.95066
0.76303 | -33.8914 | -0.97278
0.76914 | -33.8931 | -0.99343
0.77524 | -33.894 | -1.0133
0.78135 | -33.8952 | -1.0336
0.78745 | -33.8965 | -1.0537
0.79355 | -33.8976 | -1.0737
0.79966 | -33.8993 | -1.0934
0.80576 | -33.9015 | -1.1134
0.81187 | -33.9031 | -1.1352
0.81797 | -33.9049 | -1.1567
0.82408 | -33.9065 | -1.1775
0.83018 | -33.9082 | -1.1983
0.83628 | -33.91 | -1.2171
0.84239 | -33.9112 | -1.2379
0.84849 | -33.9125 | -1.2599
0.8546 | -33.9139 | -1.2801
0.8607 | -33.9151 | -1.3011
0.86681 | -33.9162 | -1.3204
0.87291 | -33.9171 | -1.3395
0.87901 | -33.9175 | -1.3579
0.88512 | -33.9182 | -1.3764
0.89122 | -33.9189 | -1.395
0.89733 | -33.9192 | -1.4111
0.90343 | -33.9196 | -1.4291
0.90953 | -33.9197 | -1.444
0.91564 | -33.9197 | -1.4604
0.92174 | -33.9199 | -1.4759
0.92785 | -33.9199 | -1.491
0.93395 | -33.92 | -1.5054
0.94006 | -33.9201 | -1.5188
0.94616 | -33.9198 | -1.5314
0.95226 | -33.9199 | -1.5447
0.95837 | -33.9199 | -1.5591
0.96447 | -33.9201 | -1.5726
0.97058 | -33.9199 | -1.5865
0.97668 | -33.9196 | -1.6006
0.98279 | -33.9193 | -1.6139
0.98889 | -33.9191 | -1.6252
0.99499 | -33.919 | -1.6393
1.0011 | -33.9188 | -1.6506
1.0072 | -33.9184 | -1.6628
1.0133 | -33.9181 | -1.6741
1.0194 | -33.9176 | -1.6877
1.0255 | -33.9173 | -1.7019
1.0316 | -33.9167 | -1.7144
1.0377 | -33.9163 | -1.7286
1.0438 | -33.9157 | -1.7418
1.0499 | -33.9149 | -1.7575
1.056 | -33.9142 | -1.7723
1.0621 | -33.9136 | -1.7871
1.0682 | -33.9131 | -1.8
1.0743 | -33.9124 | -1.8158
1.0805 | -33.9119 | -1.8306
1.0866 | -33.911 | -1.846
1.0927 | -33.9101 | -1.8608
1.0988 | -33.9095 | -1.8761
1.1049 | -33.9087 | -1.8907
1.111 | -33.9079 | -1.9064
1.1171 | -33.9072 | -1.9214
1.1232 | -33.9067 | -1.9368
1.1293 | -33.9062 | -1.9512
1.1354 | -33.9059 | -1.9665
1.1415 | -33.9056 | -1.9823
1.1476 | -33.9049 | -1.9982
1.1537 | -33.9044 | -2.0132
1.1598 | -33.9045 | -2.0282
1.1659 | -33.9042 | -2.0418
1.172 | -33.904 | -2.0561
1.1781 | -33.9037 | -2.0703
1.1842 | -33.9036 | -2.0847
1.1903 | -33.9039 | -2.096
1.1964 | -33.9038 | -2.1097
1.2025 | -33.9039 | -2.1239
1.2086 | -33.9042 | -2.1366
1.2147 | -33.9042 | -2.1493
1.2209 | -33.9046 | -2.1608
1.227 | -33.9053 | -2.1719
1.2331 | -33.9057 | -2.1842
1.2392 | -33.906 | -2.1963
1.2453 | -33.9064 | -2.2077
1.2514 | -33.9068 | -2.218
1.2575 | -33.9074 | -2.2293
1.2636 | -33.9082 | -2.2387
1.2697 | -33.9087 | -2.2483
1.2758 | -33.9094 | -2.2598
1.2819 | -33.9101 | -2.2701
1.288 | -33.9107 | -2.2812
1.2941 | -33.9115 | -2.29
1.3002 | -33.9121 | -2.2984
1.3063 | -33.9127 | -2.3078
1.3124 | -33.9133 | -2.3172
1.3185 | -33.914 | -2.3267
1.3246 | -33.9146 | -2.3361
1.3307 | -33.9154 | -2.3451
1.3368 | -33.9162 | -2.354
1.3429 | -33.9167 | -2.3633
1.349 | -33.9172 | -2.3729
1.3551 | -33.9179 | -2.3826
1.3613 | -33.9184 | -2.3913
1.3674 | -33.9189 | -2.4016
1.3735 | -33.9193 | -2.4116
1.3796 | -33.9197 | -2.4209
1.3857 | -33.92 | -2.4306
1.3918 | -33.9205 | -2.441
1.3979 | -33.921 | -2.4504
1.404 | -33.9215 | -2.462
1.4101 | -33.9218 | -2.4726
1.4162 | -33.9222 | -2.4835
1.4223 | -33.9226 | -2.4951
1.4284 | -33.9231 | -2.5054
1.4345 | -33.9233 | -2.5171
1.4406 | -33.9235 | -2.5308
1.4467 | -33.924 | -2.5427
1.4528 | -33.9244 | -2.5544
1.4589 | -33.9248 | -2.5664
1.465 | -33.925 | -2.5775
1.4711 | -33.9255 | -2.5891
1.4772 | -33.926 | -2.6027
1.4833 | -33.9264 | -2.6159
1.4894 | -33.9269 | -2.6284
1.4955 | -33.9274 | -2.6407
1.5016 | -33.9279 | -2.6535
1.5078 | -33.9282 | -2.6677
1.5139 | -33.9286 | -2.6811
1.52 | -33.929 | -2.6941
1.5261 | -33.9293 | -2.7071
1.5322 | -33.9296 | -2.7214
1.5383 | -33.93 | -2.7343
1.5444 | -33.9305 | -2.7473
1.5505 | -33.931 | -2.761
1.5566 | -33.9315 | -2.7744
1.5627 | -33.9319 | -2.7885
1.5688 | -33.9325 | -2.8025
1.5749 | -33.9328 | -2.8155
1.581 | -33.9331 | -2.8292
1.5871 | -33.9335 | -2.8433
1.5932 | -33.9338 | -2.8565
1.5993 | -33.9342 | -2.8711
1.6054 | -33.9346 | -2.8847
1.6115 | -33.935 | -2.8991
1.6176 | -33.9353 | -2.9121
1.6237 | -33.9355 | -2.9261
1.6298 | -33.9357 | -2.939
1.6359 | -33.9361 | -2.9521
1.642 | -33.9365 | -2.9643
1.6482 | -33.9366 | -2.9776
1.6543 | -33.9369 | -2.9915
1.6604 | -33.9372 | -3.004
1.6665 | -33.9373 | -3.0168
1.6726 | -33.9375 | -3.0297
1.6787 | -33.9378 | -3.0425
1.6848 | -33.9377 | -3.0553
1.6909 | -33.9379 | -3.0683
1.697 | -33.9381 | -3.0808
1.7031 | -33.9382 | -3.0937
1.7092 | -33.9381 | -3.1061
1.7153 | -33.9382 | -3.1184
1.7214 | -33.938 | -3.1309
1.7275 | -33.9379 | -3.1438
1.7336 | -33.9378 | -3.157
1.7397 | -33.938 | -3.1696
1.7458 | -33.9379 | -3.1821
1.7519 | -33.9379 | -3.1945
1.758 | -33.9379 | -3.2067
1.7641 | -33.938 | -3.2186
1.7702 | -33.9379 | -3.2318
1.7763 | -33.9379 | -3.2435
1.7824 | -33.9378 | -3.2562
1.7885 | -33.9378 | -3.269
1.7947 | -33.9379 | -3.2813
1.8008 | -33.9378 | -3.2928
1.8069 | -33.9378 | -3.3046
1.813 | -33.9379 | -3.317
1.8191 | -33.9378 | -3.33
1.8252 | -33.9379 | -3.343
1.8313 | -33.938 | -3.3555
1.8374 | -33.9381 | -3.3677
1.8435 | -33.9381 | -3.3799
1.8496 | -33.9381 | -3.3933
1.8557 | -33.9383 | -3.4064
1.8618 | -33.9385 | -3.419
1.8679 | -33.9386 | -3.4314
1.874 | -33.9387 | -3.4437
1.8801 | -33.9388 | -3.4564
1.8862 | -33.939 | -3.4687
1.8923 | -33.9391 | -3.4812
1.8984 | -33.9392 | -3.4936
1.9045 | -33.9393 | -3.5056
1.9106 | -33.9393 | -3.5182
1.9167 | -33.9393 | -3.5308
1.9228 | -33.9394 | -3.5429
1.9289 | -33.9395 | -3.5552
1.9351 | -33.9397 | -3.5664
1.9412 | -33.9399 | -3.578
1.9473 | -33.9401 | -3.5904
1.9534 | -33.9403 | -3.6021
1.9595 | -33.9405 | -3.6142
1.9656 | -33.9407 | -3.627
1.9717 | -33.9409 | -3.639
1.9778 | -33.9412 | -3.6508
1.9839 | -33.9414 | -3.6626
1.99 | -33.9416 | -3.674
1.9961 | -33.9417 | -3.6865
2.0022 | -33.9419 | -3.6984
2.0083 | -33.942 | -3.71
2.0144 | -33.9422 | -3.7218
2.0205 | -33.9425 | -3.734
2.0266 | -33.9428 | -3.7458
2.0327 | -33.943 | -3.7578
2.0388 | -33.9434 | -3.7698
2.0449 | -33.9436 | -3.7816
2.051 | -33.9438 | -3.7934
2.0571 | -33.9441 | -3.8052
2.0632 | -33.9443 | -3.817
2.0693 | -33.9445 | -3.8295
2.0754 | -33.9448 | -3.8412
2.0816 | -33.9451 | -3.8537
2.0877 | -33.9453 | -3.8656
2.0938 | -33.9456 | -3.8781
2.0999 | -33.9458 | -3.8902
2.106 | -33.946 | -3.9023
2.1121 | -33.9462 | -3.9146
2.1182 | -33.9464 | -3.9258
2.1243 | -33.9467 | -3.937
2.1304 | -33.9469 | -3.9486
2.1365 | -33.9472 | -3.9607
2.1426 | -33.9474 | -3.9732
2.1487 | -33.9477 | -3.9849
2.1548 | -33.9481 | -3.9971
2.1609 | -33.9486 | -4.009
2.167 | -33.9489 | -4.0206
2.1731 | -33.9492 | -4.0331
2.1792 | -33.9495 | -4.0446
2.1853 | -33.9498 | -4.0569
2.1914 | -33.9501 | -4.0687
2.1975 | -33.9505 | -4.0807
2.2036 | -33.9507 | -4.0932
2.2097 | -33.951 | -4.1052
2.2158 | -33.9514 | -4.1163
2.222 | -33.9517 | -4.1273
2.2281 | -33.9519 | -4.1391
2.2342 | -33.9522 | -4.1509
2.2403 | -33.9524 | -4.162
2.2464 | -33.9527 | -4.1732
2.2525 | -33.953 | -4.1851
2.2586 | -33.9533 | -4.1961
2.2647 | -33.9536 | -4.2071
2.2708 | -33.9539 | -4.2188
2.2769 | -33.9543 | -4.2299
2.283 | -33.9545 | -4.2415
2.2891 | -33.9547 | -4.2528
2.2952 | -33.9549 | -4.2638
2.3013 | -33.9551 | -4.2752
2.3074 | -33.9554 | -4.2861
2.3135 | -33.9556 | -4.2979
2.3196 | -33.9559 | -4.3086
2.3257 | -33.9561 | -4.3197
2.3318 | -33.9562 | -4.3312
2.3379 | -33.9564 | -4.3425
2.344 | -33.9566 | -4.3543
2.3501 | -33.9568 | -4.3655
2.3562 | -33.9569 | -4.3763
2.3623 | -33.9571 | -4.3871
2.3685 | -33.9572 | -4.3981
2.3746 | -33.9573 | -4.4097
2.3807 | -33.9575 | -4.4207
2.3868 | -33.9577 | -4.4318
2.3929 | -33.9579 | -4.4429
2.399 | -33.958 | -4.4538
2.4051 | -33.9583 | -4.4648
2.4112 | -33.9585 | -4.4759
2.4173 | -33.9586 | -4.4869
2.4234 | -33.9587 | -4.4981
2.4295 | -33.9588 | -4.5093
2.4356 | -33.959 | -4.5206
2.4417 | -33.9591 | -4.5322
2.4478 | -33.9592 | -4.5434
2.4539 | -33.9594 | -4.554
2.46 | -33.9595 | -4.5658
2.4661 | -33.9596 | -4.5772
2.4722 | -33.9598 | -4.5878
2.4783 | -33.9599 | -4.5986
2.4844 | -33.96 | -4.61
2.4905 | -33.96 | -4.6213
2.4966 | -33.9602 | -4.6324
2.5027 | -33.9603 | -4.6432
2.5089 | -33.9604 | -4.6544
2.515 | -33.9604 | -4.6656
2.5211 | -33.9605 | -4.6767
2.5272 | -33.9606 | -4.688
2.5333 | -33.9607 | -4.6993
2.5394 | -33.9608 | -4.7104
2.5455 | -33.9609 | -4.7221
2.5516 | -33.961 | -4.7328
2.5577 | -33.9612 | -4.7445
2.5638 | -33.9614 | -4.7564
2.5699 | -33.9614 | -4.7681
2.576 | -33.9614 | -4.78
2.5821 | -33.9615 | -4.7915
2.5882 | -33.9616 | -4.803
2.5943 | -33.9617 | -4.8141
2.6004 | -33.9619 | -4.8253
2.6065 | -33.962 | -4.836
2.6126 | -33.9621 | -4.8467
2.6187 | -33.9622 | -4.8575
2.6248 | -33.9624 | -4.8685
2.6309 | -33.9625 | -4.8802
2.637 | -33.9627 | -4.8921
2.6431 | -33.9629 | -4.9028
2.6492 | -33.963 | -4.9144
2.6554 | -33.9632 | -4.9255
2.6615 | -33.9632 | -4.9367
2.6676 | -33.9635 | -4.9483
2.6737 | -33.9638 | -4.9597
2.6798 | -33.964 | -4.971
2.6859 | -33.9642 | -4.9826
2.692 | -33.9643 | -4.9937
2.6981 | -33.9645 | -5.005
2.7042 | -33.9647 | -5.0164
2.7103 | -33.9649 | -5.0274
2.7164 | -33.9652 | -5.0386
2.7225 | -33.9654 | -5.0501
2.7286 | -33.9657 | -5.0614
2.7347 | -33.9658 | -5.0722
2.7408 | -33.9661 | -5.0826
2.7469 | -33.9663 | -5.0941
2.753 | -33.9666 | -5.1049
2.7591 | -33.9669 | -5.1165
2.7652 | -33.9671 | -5.1282
2.7713 | -33.9673 | -5.1391
2.7774 | -33.9676 | -5.1503
2.7835 | -33.9677 | -5.1615
2.7896 | -33.968 | -5.1723
2.7958 | -33.9682 | -5.1834
2.8019 | -33.9685 | -5.194
2.808 | -33.9688 | -5.2049
2.8141 | -33.9691 | -5.2163
2.8202 | -33.9694 | -5.2274
2.8263 | -33.9697 | -5.2386
2.8324 | -33.9701 | -5.2502
2.8385 | -33.9705 | -5.2613
2.8446 | -33.9709 | -5.2723
2.8507 | -33.9712 | -5.2834
2.8568 | -33.9716 | -5.294
2.8629 | -33.9718 | -5.3049
2.869 | -33.9722 | -5.3163
2.8751 | -33.9725 | -5.3273
2.8812 | -33.9729 | -5.3389
2.8873 | -33.9732 | -5.35
2.8934 | -33.9735 | -5.3613
2.8995 | -33.9739 | -5.3726
2.9056 | -33.9742 | -5.384
2.9117 | -33.9745 | -5.3954
2.9178 | -33.9749 | -5.4064
2.9239 | -33.9753 | -5.4182
2.93 | -33.9757 | -5.43
2.9361 | -33.976 | -5.4411
2.9423 | -33.9763 | -5.4521
2.9484 | -33.9768 | -5.4635
2.9545 | -33.9772 | -5.4749
2.9606 | -33.9776 | -5.4866
2.9667 | -33.9779 | -5.498
2.9728 | -33.9782 | -5.5097
2.9789 | -33.9787 | -5.5211
2.985 | -33.9791 | -5.5327
2.9911 | -33.9795 | -5.5443
2.9972 | -33.9799 | -5.5557
3.0033 | -33.9803 | -5.5675
3.0094 | -33.9807 | -5.5789
3.0155 | -33.9812 | -5.5902
3.0216 | -33.9815 | -5.6018
3.0277 | -33.9818 | -5.6137
3.0338 | -33.9822 | -5.6255
3.0399 | -33.9826 | -5.6372
3.046 | -33.9829 | -5.6491
3.0521 | -33.9833 | -5.6605
3.0582 | -33.9836 | -5.673
3.0643 | -33.9839 | -5.6848
3.0704 | -33.9842 | -5.697
3.0765 | -33.9845 | -5.7084
3.0827 | -33.9848 | -5.7199
3.0888 | -33.9851 | -5.7316
3.0949 | -33.9855 | -5.7434
3.101 | -33.9859 | -5.7552
3.1071 | -33.9862 | -5.7674
3.1132 | -33.9865 | -5.7795
3.1193 | -33.9869 | -5.7914
3.1254 | -33.9873 | -5.8032
3.1315 | -33.9876 | -5.8155
3.1376 | -33.9879 | -5.8273
3.1437 | -33.9882 | -5.8392
3.1498 | -33.9886 | -5.8517
3.1559 | -33.9889 | -5.8634
3.162 | -33.9891 | -5.8759
3.1681 | -33.9894 | -5.8882
3.1742 | -33.9897 | -5.9003
3.1803 | -33.99 | -5.9123
3.1864 | -33.9904 | -5.924
3.1925 | -33.9906 | -5.9359
3.1986 | -33.9909 | -5.9484
3.2047 | -33.9912 | -5.9609
3.2108 | -33.9914 | -5.9724
3.2169 | -33.9917 | -5.9845
3.223 | -33.9919 | -5.9966
3.2292 | -33.9922 | -6.0088
3.2353 | -33.9924 | -6.0201
3.2414 | -33.9926 | -6.0319
3.2475 | -33.9928 | -6.0438
3.2536 | -33.9931 | -6.0557
3.2597 | -33.9933 | -6.0671
3.2658 | -33.9935 | -6.0785
3.2719 | -33.9937 | -6.09
3.278 | -33.994 | -6.1019
3.2841 | -33.9942 | -6.1134
3.2902 | -33.9945 | -6.1253
3.2963 | -33.9947 | -6.1373
3.3024 | -33.9948 | -6.1487
3.3085 | -33.9951 | -6.1606
3.3146 | -33.9952 | -6.1722
3.3207 | -33.9955 | -6.184
3.3268 | -33.9957 | -6.1959
3.3329 | -33.9959 | -6.2071
3.339 | -33.996 | -6.2183
3.3451 | -33.9962 | -6.23
3.3512 | -33.9964 | -6.2411
3.3573 | -33.9966 | -6.2527
3.3634 | -33.9968 | -6.2643
3.3696 | -33.997 | -6.2761
3.3757 | -33.9971 | -6.2873
3.3818 | -33.9974 | -6.2986
3.3879 | -33.9975 | -6.3101
3.394 | -33.9977 | -6.3218
3.4001 | -33.9979 | -6.333
3.4062 | -33.9979 | -6.3444
3.4123 | -33.9981 | -6.3553
3.4184 | -33.9983 | -6.367
3.4245 | -33.9985 | -6.3782
3.4306 | -33.9986 | -6.3893
3.4367 | -33.9988 | -6.4002
3.4428 | -33.9989 | -6.4109
3.4489 | -33.9991 | -6.4223
3.455 | -33.9992 | -6.4335
3.4611 | -33.9994 | -6.4451
3.4672 | -33.9996 | -6.4566
3.4733 | -33.9997 | -6.468
3.4794 | -33.9999 | -6.4792
3.4855 | -34 | -6.4899
3.4916 | -34.0002 | -6.5007
3.4977 | -34.0004 | -6.5124
3.5038 | -34.0005 | -6.5235
3.5099 | -34.0006 | -6.5342
3.5161 | -34.0008 | -6.5449
3.5222 | -34.0009 | -6.5556
3.5283 | -34.001 | -6.5668
3.5344 | -34.0012 | -6.5779
3.5405 | -34.0014 | -6.5889
3.5466 | -34.0015 | -6.6002
3.5527 | -34.0016 | -6.6113
3.5588 | -34.0018 | -6.6223
3.5649 | -34.0018 | -6.6332
3.571 | -34.002 | -6.6442
3.5771 | -34.0021 | -6.6556
3.5832 | -34.0023 | -6.6658
3.5893 | -34.0025 | -6.6768
3.5954 | -34.0027 | -6.6879
3.6015 | -34.0028 | -6.6991
3.6076 | -34.0031 | -6.71
3.6137 | -34.0032 | -6.7204
3.6198 | -34.0035 | -6.7309
3.6259 | -34.0036 | -6.7422
3.632 | -34.0038 | -6.7531
3.6381 | -34.004 | -6.7641
3.6442 | -34.0042 | -6.775
3.6503 | -34.0044 | -6.7854
3.6565 | -34.0046 | -6.7965
3.6626 | -34.0047 | -6.8074
3.6687 | -34.005 | -6.8185
3.6748 | -34.0053 | -6.8292
3.6809 | -34.0055 | -6.8401
3.687 | -34.0057 | -6.8514
3.6931 | -34.006 | -6.8618
3.6992 | -34.0062 | -6.8728
3.7053 | -34.0064 | -6.8836
3.7114 | -34.0065 | -6.8946
3.7175 | -34.0068 | -6.9055
3.7236 | -34.0071 | -6.9159
3.7297 | -34.0073 | -6.9269
3.7358 | -34.0075 | -6.9381
3.7419 | -34.0077 | -6.9493
3.748 | -34.0079 | -6.9602
3.7541 | -34.0082 | -6.9711
3.7602 | -34.0085 | -6.9818
3.7663 | -34.0087 | -6.9926
3.7724 | -34.0089 | -7.0035
3.7785 | -34.0092 | -7.0142
3.7846 | -34.0094 | -7.0249
3.7907 | -34.0098 | -7.0361
3.7969 | -34.01 | -7.0471
3.803 | -34.0103 | -7.0577
3.8091 | -34.0105 | -7.0685
3.8152 | -34.0108 | -7.0792
3.8213 | -34.0111 | -7.0904
3.8274 | -34.0114 | -7.1009
3.8335 | -34.0117 | -7.112
3.8396 | -34.012 | -7.1228
3.8457 | -34.0124 | -7.1338
3.8518 | -34.0127 | -7.1447
3.8579 | -34.0131 | -7.1555
3.864 | -34.0134 | -7.1666
3.8701 | -34.0137 | -7.1779
3.8762 | -34.014 | -7.1886
3.8823 | -34.0142 | -7.1998
3.8884 | -34.0146 | -7.2107
3.8945 | -34.015 | -7.2219
3.9006 | -34.0153 | -7.2332
3.9067 | -34.0156 | -7.2446
3.9128 | -34.0159 | -7.2557
3.9189 | -34.0162 | -7.2664
3.925 | -34.0165 | -7.2776
3.9311 | -34.0168 | -7.2881
3.9372 | -34.0171 | -7.2994
3.9434 | -34.0174 | -7.3105
3.9495 | -34.0178 | -7.3216
3.9556 | -34.0181 | -7.3324
3.9617 | -34.0184 | -7.3431
3.9678 | -34.0188 | -7.3538
3.9739 | -34.0191 | -7.3648
3.98 | -34.0194 | -7.376
3.9861 | -34.0197 | -7.3869
3.9922 | -34.0199 | -7.3979
3.9983 | -34.0203 | -7.4086
4.0044 | -34.0206 | -7.4192
4.0105 | -34.0209 | -7.4302
4.0166 | -34.0212 | -7.4413
4.0227 | -34.0216 | -7.4523
4.0288 | -34.0218 | -7.4637
4.0349 | -34.0221 | -7.4744
4.041 | -34.0224 | -7.4854
4.0471 | -34.0226 | -7.4965
4.0532 | -34.0229 | -7.5071
4.0593 | -34.0233 | -7.5182
4.0654 | -34.0236 | -7.529
4.0715 | -34.0239 | -7.5404
4.0776 | -34.0242 | -7.5513
4.0838 | -34.0245 | -7.5627
4.0899 | -34.0248 | -7.5734
4.096 | -34.0251 | -7.5844
4.1021 | -34.0254 | -7.5951
4.1082 | -34.0257 | -7.6058
4.1143 | -34.026 | -7.6165
4.1204 | -34.0263 | -7.6272
4.1265 | -34.0266 | -7.6378
4.1326 | -34.0269 | -7.6485
4.1387 | -34.0272 | -7.6593
4.1448 | -34.0275 | -7.6703
4.1509 | -34.0277 | -7.6813
4.157 | -34.028 | -7.6927
4.1631 | -34.0283 | -7.7038
4.1692 | -34.0286 | -7.7145
4.1753 | -34.0289 | -7.7254
4.1814 | -34.0291 | -7.736
4.1875 | -34.0294 | -7.7472
4.1936 | -34.0297 | -7.7583
4.1997 | -34.0299 | -7.7688
4.2058 | -34.0302 | -7.78
4.2119 | -34.0305 | -7.7908
4.218 | -34.0307 | -7.8014
4.2241 | -34.031 | -7.8121
4.2303 | -34.0313 | -7.8229
4.2364 | -34.0315 | -7.8337
4.2425 | -34.0317 | -7.8446
4.2486 | -34.032 | -7.8556
4.2547 | -34.0321 | -7.8662
4.2608 | -34.0324 | -7.8772
4.2669 | -34.0326 | -7.8878
4.273 | -34.0329 | -7.8988
4.2791 | -34.0331 | -7.9097
4.2852 | -34.0334 | -7.9207
4.2913 | -34.0335 | -7.9313
4.2974 | -34.0338 | -7.9424
4.3035 | -34.0341 | -7.9535
4.3096 | -34.0343 | -7.9646
4.3157 | -34.0346 | -7.9753
4.3218 | -34.0348 | -7.9861
4.3279 | -34.035 | -7.9969
4.334 | -34.0352 | -8.0076
4.3401 | -34.0355 | -8.0181
4.3462 | -34.0358 | -8.029
4.3523 | -34.036 | -8.0394
4.3584 | -34.0362 | -8.0497
4.3645 | -34.0364 | -8.0604
4.3707 | -34.0367 | -8.0709
4.3768 | -34.0369 | -8.0815
4.3829 | -34.0372 | -8.0926
4.389 | -34.0374 | -8.1034
4.3951 | -34.0376 | -8.1139
4.4012 | -34.0378 | -8.1249
4.4073 | -34.038 | -8.1355
4.4134 | -34.0382 | -8.1461
4.4195 | -34.0385 | -8.1568
4.4256 | -34.0388 | -8.1673
4.4317 | -34.039 | -8.1779
4.4378 | -34.0392 | -8.1883
4.4439 | -34.0395 | -8.1988
4.45 | -34.0396 | -8.2094
4.4561 | -34.0398 | -8.2201
4.4622 | -34.04 | -8.2308
4.4683 | -34.0403 | -8.2414
4.4744 | -34.0405 | -8.2524
4.4805 | -34.0407 | -8.2627
4.4866 | -34.0409 | -8.2733
4.4927 | -34.0412 | -8.2842
4.4988 | -34.0414 | -8.2949
4.5049 | -34.0417 | -8.3051
4.511 | -34.0419 | -8.3154
4.5172 | -34.0422 | -8.326
4.5233 | -34.0423 | -8.337
4.5294 | -34.0426 | -8.3475
4.5355 | -34.0428 | -8.3581
4.5416 | -34.043 | -8.3686
4.5477 | -34.0433 | -8.3793
4.5538 | -34.0435 | -8.3894
4.5599 | -34.0438 | -8.4
4.566 | -34.044 | -8.4106
4.5721 | -34.0442 | -8.4212
4.5782 | -34.0445 | -8.432
4.5843 | -34.0447 | -8.4427
4.5904 | -34.045 | -8.4534
4.5965 | -34.0453 | -8.4638
4.6026 | -34.0455 | -8.4749
4.6087 | -34.0457 | -8.4854
4.6148 | -34.046 | -8.4964
4.6209 | -34.0462 | -8.5069
4.627 | -34.0465 | -8.5176
4.6331 | -34.0467 | -8.5281
4.6392 | -34.0469 | -8.5388
4.6453 | -34.0471 | -8.5494
4.6514 | -34.0474 | -8.56
4.6576 | -34.0476 | -8.571
4.6637 | -34.0478 | -8.5812
4.6698 | -34.048 | -8.5917
4.6759 | -34.0483 | -8.6027
4.682 | -34.0485 | -8.6132
4.6881 | -34.0487 | -8.6237
4.6942 | -34.049 | -8.6343
4.7003 | -34.0492 | -8.6449
4.7064 | -34.0494 | -8.6553
4.7125 | -34.0497 | -8.6661
4.7186 | -34.0499 | -8.6764
4.7247 | -34.0502 | -8.6871
4.7308 | -34.0505 | -8.6976
4.7369 | -34.0507 | -8.7083
4.743 | -34.051 | -8.7193
4.7491 | -34.0512 | -8.7302
4.7552 | -34.0514 | -8.7409
4.7613 | -34.0516 | -8.7517
4.7674 | -34.0519 | -8.7621
4.7735 | -34.0521 | -8.7732
4.7796 | -34.0524 | -8.784
4.7857 | -34.0526 | -8.7947
4.7918 | -34.0529 | -8.8051
4.7979 | -34.0532 | -8.8156
4.8041 | -34.0534 | -8.8263
4.8102 | -34.0537 | -8.8366
4.8163 | -34.0539 | -8.8471
4.8224 | -34.0542 | -8.8578
4.8285 | -34.0544 | -8.8681
4.8346 | -34.0547 | -8.8786
4.8407 | -34.0549 | -8.8893
4.8468 | -34.0552 | -8.9001
4.8529 | -34.0555 | -8.911
4.859 | -34.0558 | -8.9219
4.8651 | -34.056 | -8.9323
4.8712 | -34.0563 | -8.9427
4.8773 | -34.0565 | -8.9534
4.8834 | -34.0568 | -8.964
4.8895 | -34.057 | -8.9743
4.8956 | -34.0572 | -8.985
4.9017 | -34.0574 | -8.9955
4.9078 | -34.0577 | -9.0064
4.9139 | -34.058 | -9.0166
4.92 | -34.0583 | -9.0269
4.9261 | -34.0586 | -9.0371
4.9322 | -34.0588 | -9.0477
4.9383 | -34.0591 | -9.0583
4.9445 | -34.0594 | -9.0689
4.9506 | -34.0596 | -9.0793
4.9567 | -34.0599 | -9.0897
4.9628 | -34.0602 | -9.1002
4.9689 | -34.0604 | -9.1112
4.975 | -34.0607 | -9.122
4.9811 | -34.061 | -9.1323
4.9872 | -34.0613 | -9.1429
4.9933 | -34.0616 | -9.1537
4.9994 | -34.0619 | -9.1647
5.0055 | -34.0622 | -9.1753
5.0116 | -34.0625 | -9.1863
5.0177 | -34.0628 | -9.1972
5.0238 | -34.0631 | -9.2081
5.0299 | -34.0634 | -9.219
5.036 | -34.0637 | -9.2297
5.0421 | -34.064 | -9.2402
5.0482 | -34.0643 | -9.2509
5.0543 | -34.0646 | -9.2613
5.0604 | -34.0649 | -9.2719
5.0665 | -34.0652 | -9.2822
5.0726 | -34.0655 | -9.2932
5.0787 | -34.0658 | -9.3041
5.0848 | -34.0661 | -9.3149
5.091 | -34.0663 | -9.3253
5.0971 | -34.0667 | -9.3356
5.1032 | -34.067 | -9.3466
5.1093 | -34.0673 | -9.3572
5.1154 | -34.0675 | -9.3676
5.1215 | -34.0678 | -9.378
5.1276 | -34.0681 | -9.3884
5.1337 | -34.0685 | -9.3987
5.1398 | -34.0687 | -9.4091
5.1459 | -34.0691 | -9.4199
5.152 | -34.0694 | -9.4307
5.1581 | -34.0696 | -9.4415
5.1642 | -34.07 | -9.4523
5.1703 | -34.0703 | -9.4629
5.1764 | -34.0706 | -9.4734
5.1825 | -34.0709 | -9.4838
5.1886 | -34.0712 | -9.4942
5.1947 | -34.0715 | -9.5049
5.2008 | -34.0718 | -9.5156
5.2069 | -34.0721 | -9.5264
5.213 | -34.0724 | -9.5368
5.2191 | -34.0728 | -9.5476
5.2252 | -34.0731 | -9.5582
5.2314 | -34.0733 | -9.5691
5.2375 | -34.0737 | -9.5794
5.2436 | -34.074 | -9.5902
5.2497 | -34.0743 | -9.6006
5.2558 | -34.0746 | -9.6116
5.2619 | -34.0749 | -9.6224
5.268 | -34.0753 | -9.6332
5.2741 | -34.0756 | -9.644
5.2802 | -34.0759 | -9.655
5.2863 | -34.0763 | -9.6655
5.2924 | -34.0766 | -9.6764
5.2985 | -34.0769 | -9.6871
5.3046 | -34.0772 | -9.6978
5.3107 | -34.0776 | -9.7088
5.3168 | -34.0779 | -9.7196
5.3229 | -34.0782 | -9.7303
5.329 | -34.0785 | -9.741
5.3351 | -34.0789 | -9.7514
5.3412 | -34.0792 | -9.7621
5.3473 | -34.0795 | -9.7729
5.3534 | -34.0798 | -9.7837
5.3595 | -34.0801 | -9.7944
5.3656 | -34.0805 | -9.8051
5.3717 | -34.0808 | -9.8162
5.3779 | -34.0811 | -9.8272
5.384 | -34.0814 | -9.8379
5.3901 | -34.0818 | -9.8489
5.3962 | -34.0821 | -9.8595
5.4023 | -34.0824 | -9.8705
5.4084 | -34.0828 | -9.8811
5.4145 | -34.0832 | -9.8917
5.4206 | -34.0835 | -9.9025
5.4267 | -34.0838 | -9.9135
5.4328 | -34.0841 | -9.9243
5.4389 | -34.0845 | -9.9348
5.445 | -34.0847 | -9.946
5.4511 | -34.0851 | -9.9565
5.4572 | -34.0854 | -9.9673
5.4633 | -34.0858 | -9.978
5.4694 | -34.0861 | -9.9889
5.4755 | -34.0864 | -9.9995
5.4816 | -34.0867 | -10.0103
5.4877 | -34.0871 | -10.0209
5.4938 | -34.0874 | -10.0314
5.4999 | -34.0877 | -10.0421
5.506 | -34.0881 | -10.0529
5.5121 | -34.0884 | -10.0635
5.5183 | -34.0888 | -10.0741
5.5244 | -34.0891 | -10.0851
5.5305 | -34.0894 | -10.0957
5.5366 | -34.0897 | -10.1063
5.5427 | -34.09 | -10.1172
5.5488 | -34.0904 | -10.1279
5.5549 | -34.0907 | -10.1388
5.561 | -34.091 | -10.1498
5.5671 | -34.0914 | -10.1605
5.5732 | -34.0917 | -10.1715
5.5793 | -34.092 | -10.1826
5.5854 | -34.0924 | -10.1933
5.5915 | -34.0927 | -10.2045
5.5976 | -34.093 | -10.2154
5.6037 | -34.0934 | -10.2262
5.6098 | -34.0937 | -10.2369
5.6159 | -34.094 | -10.2475
5.622 | -34.0943 | -10.2583
5.6281 | -34.0947 | -10.2693
5.6342 | -34.095 | -10.2804
5.6403 | -34.0953 | -10.2913
5.6464 | -34.0957 | -10.3021
5.6525 | -34.096 | -10.3131
5.6586 | -34.0964 | -10.3236
5.6648 | -34.0967 | -10.3345
5.6709 | -34.0971 | -10.3455
5.677 | -34.0974 | -10.3563
5.6831 | -34.0978 | -10.3672
5.6892 | -34.098 | -10.3781
5.6953 | -34.0984 | -10.3889
5.7014 | -34.0987 | -10.3999
5.7075 | -34.099 | -10.4107
5.7136 | -34.0994 | -10.4214
5.7197 | -34.0997 | -10.4321
5.7258 | -34.1 | -10.4428
5.7319 | -34.1004 | -10.4537
5.738 | -34.1007 | -10.4645
5.7441 | -34.1011 | -10.4753
5.7502 | -34.1014 | -10.4858
5.7563 | -34.1018 | -10.4967
5.7624 | -34.1021 | -10.5077
5.7685 | -34.1025 | -10.5185
5.7746 | -34.1028 | -10.5297
5.7807 | -34.103 | -10.5406
5.7868 | -34.1034 | -10.5516
5.7929 | -34.1037 | -10.5623
5.799 | -34.104 | -10.5731
5.8052 | -34.1043 | -10.5837
5.8113 | -34.1046 | -10.5944
5.8174 | -34.1049 | -10.6051
5.8235 | -34.1053 | -10.6158
5.8296 | -34.1056 | -10.6268
5.8357 | -34.1059 | -10.6375
5.8418 | -34.1062 | -10.6484
5.8479 | -34.1065 | -10.6593
5.854 | -34.1068 | -10.6702
5.8601 | -34.1072 | -10.6809
5.8662 | -34.1075 | -10.6919
5.8723 | -34.1078 | -10.7025
5.8784 | -34.1081 | -10.7132
5.8845 | -34.1084 | -10.7239
5.8906 | -34.1087 | -10.7347
5.8967 | -34.109 | -10.7454
5.9028 | -34.1094 | -10.756
5.9089 | -34.1098 | -10.7666
5.915 | -34.1101 | -10.7772
5.9211 | -34.1104 | -10.7878
5.9272 | -34.1107 | -10.7984
5.9333 | -34.111 | -10.809
5.9394 | -34.1114 | -10.8197
5.9455 | -34.1117 | -10.8303
5.9517 | -34.1121 | -10.8408
5.9578 | -34.1125 | -10.8516
5.9639 | -34.1128 | -10.862
5.97 | -34.1132 | -10.8728
5.9761 | -34.1135 | -10.8834
5.9822 | -34.1138 | -10.8942
5.9883 | -34.1141 | -10.9049
5.9944 | -34.1145 | -10.9158
6.0005 | -34.1148 | -10.9266
6.0066 | -34.1151 | -10.9371
6.0127 | -34.1155 | -10.9475
6.0188 | -34.1158 | -10.9585
6.0249 | -34.1161 | -10.9693
6.031 | -34.1164 | -10.9798
6.0371 | -34.1167 | -10.9905
6.0432 | -34.1171 | -11.001
6.0493 | -34.1174 | -11.0115
6.0554 | -34.1178 | -11.022
6.0615 | -34.1181 | -11.0327
6.0676 | -34.1184 | -11.0435
6.0737 | -34.1188 | -11.0543
6.0798 | -34.1191 | -11.0649
6.0859 | -34.1194 | -11.0754
6.0921 | -34.1198 | -11.0859
6.0982 | -34.1201 | -11.0966
6.1043 | -34.1204 | -11.1071
6.1104 | -34.1208 | -11.1179
6.1165 | -34.1211 | -11.1281
6.1226 | -34.1215 | -11.1387
6.1287 | -34.1218 | -11.1493
6.1348 | -34.1221 | -11.1599
6.1409 | -34.1225 | -11.1705
6.147 | -34.1228 | -11.181
6.1531 | -34.1231 | -11.1916
6.1592 | -34.1234 | -11.2021
6.1653 | -34.1237 | -11.2126
6.1714 | -34.1241 | -11.2232
6.1775 | -34.1244 | -11.2339
6.1836 | -34.1247 | -11.2445
6.1897 | -34.1251 | -11.2551
6.1958 | -34.1254 | -11.2659
6.2019 | -34.1257 | -11.2767
6.208 | -34.1261 | -11.2873
6.2141 | -34.1264 | -11.2979
6.2202 | -34.1268 | -11.3086
6.2263 | -34.1271 | -11.319
6.2325 | -34.1274 | -11.3293
6.2386 | -34.1278 | -11.3396
6.2447 | -34.1281 | -11.3504
6.2508 | -34.1284 | -11.361
6.2569 | -34.1288 | -11.3715
6.263 | -34.1291 | -11.382
6.2691 | -34.1294 | -11.3925
6.2752 | -34.1298 | -11.403
6.2813 | -34.1301 | -11.4136
6.2874 | -34.1304 | -11.424
6.2935 | -34.1307 | -11.4344
6.2996 | -34.131 | -11.4449
6.3057 | -34.1314 | -11.4555
6.3118 | -34.1317 | -11.4661
6.3179 | -34.132 | -11.4766
6.324 | -34.1323 | -11.487
6.3301 | -34.1327 | -11.4975
6.3362 | -34.133 | -11.5079
6.3423 | -34.1334 | -11.5186
6.3484 | -34.1337 | -11.5292
6.3545 | -34.134 | -11.5393
6.3606 | -34.1343 | -11.5496
6.3667 | -34.1346 | -11.5599
6.3728 | -34.135 | -11.5705
6.379 | -34.1353 | -11.581
6.3851 | -34.1356 | -11.5913
6.3912 | -34.1359 | -11.6017
6.3973 | -34.1362 | -11.6122
6.4034 | -34.1365 | -11.6224
6.4095 | -34.1369 | -11.6328
6.4156 | -34.1372 | -11.6433
6.4217 | -34.1375 | -11.6537
6.4278 | -34.1378 | -11.6643
6.4339 | -34.1382 | -11.6747
6.44 | -34.1385 | -11.685
6.4461 | -34.1388 | -11.6957
6.4522 | -34.1392 | -11.7059
6.4583 | -34.1395 | -11.7165
6.4644 | -34.1398 | -11.727
6.4705 | -34.1401 | -11.7371
6.4766 | -34.1404 | -11.7476
6.4827 | -34.1408 | -11.7581
6.4888 | -34.1411 | -11.7683
6.4949 | -34.1414 | -11.7785
6.501 | -34.1417 | -11.7889
6.5071 | -34.1421 | -11.799
6.5132 | -34.1424 | -11.8095
6.5194 | -34.1427 | -11.82
6.5255 | -34.143 | -11.8301
6.5316 | -34.1433 | -11.8407
6.5377 | -34.1436 | -11.851
6.5438 | -34.1439 | -11.8612
6.5499 | -34.1443 | -11.8715
6.556 | -34.1446 | -11.8818
6.5621 | -34.1449 | -11.8924
6.5682 | -34.1452 | -11.9028
6.5743 | -34.1456 | -11.9131
6.5804 | -34.1459 | -11.9235
6.5865 | -34.1462 | -11.934
6.5926 | -34.1466 | -11.9442
6.5987 | -34.1469 | -11.9543
6.6048 | -34.1472 | -11.9647
6.6109 | -34.1476 | -11.975
6.617 | -34.1479 | -11.9855
6.6231 | -34.1482 | -11.9958
6.6292 | -34.1485 | -12.0062
6.6353 | -34.1488 | -12.0169
6.6414 | -34.1492 | -12.0272
6.6475 | -34.1495 | -12.0375
6.6536 | -34.1498 | -12.0478
6.6597 | -34.1501 | -12.0579
6.6659 | -34.1505 | -12.0683
6.672 | -34.1508 | -12.0785
6.6781 | -34.1511 | -12.0889
6.6842 | -34.1515 | -12.099
6.6903 | -34.1519 | -12.1093
6.6964 | -34.1522 | -12.1197
6.7025 | -34.1525 | -12.1299
6.7086 | -34.1528 | -12.1402
6.7147 | -34.1532 | -12.1506
6.7208 | -34.1535 | -12.1611
6.7269 | -34.1538 | -12.1714
6.733 | -34.1541 | -12.1815
6.7391 | -34.1544 | -12.1918
6.7452 | -34.1548 | -12.2019
6.7513 | -34.1551 | -12.2119
6.7574 | -34.1555 | -12.2223
6.7635 | -34.1558 | -12.2325
6.7696 | -34.1561 | -12.2426
6.7757 | -34.1565 | -12.253
6.7818 | -34.1568 | -12.2633
6.7879 | -34.1572 | -12.2734
6.794 | -34.1575 | -12.2834
6.8001 | -34.1578 | -12.2937
6.8063 | -34.1582 | -12.3039
6.8124 | -34.1585 | -12.3143
6.8185 | -34.1588 | -12.3247
6.8246 | -34.1591 | -12.3348
6.8307 | -34.1594 | -12.345
6.8368 | -34.1598 | -12.3555
6.8429 | -34.1601 | -12.3657
6.849 | -34.1605 | -12.3761
6.8551 | -34.1609 | -12.3862
6.8612 | -34.1612 | -12.3964
6.8673 | -34.1615 | -12.4067
6.8734 | -34.1618 | -12.417
6.8795 | -34.1622 | -12.4273
6.8856 | -34.1625 | -12.4375
6.8917 | -34.1628 | -12.4477
6.8978 | -34.1631 | -12.4577
6.9039 | -34.1635 | -12.4677
6.91 | -34.1639 | -12.4779
6.9161 | -34.1642 | -12.488
6.9222 | -34.1645 | -12.4982
6.9283 | -34.1649 | -12.5081
6.9344 | -34.1653 | -12.5184
6.9405 | -34.1656 | -12.5285
6.9466 | -34.166 | -12.5389
6.9528 | -34.1663 | -12.5493
6.9589 | -34.1667 | -12.5596
6.965 | -34.167 | -12.5699
6.9711 | -34.1674 | -12.5801
6.9772 | -34.1678 | -12.5903
6.9833 | -34.1681 | -12.6005
6.9894 | -34.1685 | -12.611
6.9955 | -34.1688 | -12.6209
7.0016 | -34.1692 | -12.6311
7.0077 | -34.1695 | -12.6413
7.0138 | -34.1699 | -12.6515
7.0199 | -34.1702 | -12.6616
7.026 | -34.1705 | -12.6717
7.0321 | -34.1709 | -12.6817
7.0382 | -34.1713 | -12.6915
7.0443 | -34.1716 | -12.702
7.0504 | -34.172 | -12.7123
7.0565 | -34.1723 | -12.7224
7.0626 | -34.1727 | -12.7323
7.0687 | -34.173 | -12.7422
7.0748 | -34.1734 | -12.7521
7.0809 | -34.1738 | -12.7622
7.087 | -34.1741 | -12.7724
7.0932 | -34.1744 | -12.7825
7.0993 | -34.1748 | -12.7927
7.1054 | -34.1752 | -12.8027
7.1115 | -34.1755 | -12.8126
7.1176 | -34.1759 | -12.8227
7.1237 | -34.1762 | -12.8328
7.1298 | -34.1766 | -12.8428
7.1359 | -34.1769 | -12.853
7.142 | -34.1773 | -12.8628
7.1481 | -34.1776 | -12.8727
7.1542 | -34.178 | -12.8829
7.1603 | -34.1783 | -12.893
7.1664 | -34.1787 | -12.9031
7.1725 | -34.1791 | -12.9133
7.1786 | -34.1795 | -12.9234
7.1847 | -34.1798 | -12.9335
7.1908 | -34.1802 | -12.9438
7.1969 | -34.1805 | -12.9542
7.203 | -34.1809 | -12.9642
7.2091 | -34.1813 | -12.9741
7.2152 | -34.1816 | -12.9842
7.2213 | -34.182 | -12.9942
7.2274 | -34.1823 | -13.0042
7.2335 | -34.1827 | -13.0143
7.2397 | -34.1831 | -13.0243
7.2458 | -34.1835 | -13.0342
7.2519 | -34.1838 | -13.0443
7.258 | -34.1842 | -13.0546
7.2641 | -34.1845 | -13.0647
7.2702 | -34.1849 | -13.075
7.2763 | -34.1853 | -13.0852
7.2824 | -34.1857 | -13.0954
7.2885 | -34.186 | -13.1056
7.2946 | -34.1864 | -13.1157
7.3007 | -34.1867 | -13.1259
7.3068 | -34.187 | -13.1362
7.3129 | -34.1874 | -13.1464
7.319 | -34.1878 | -13.1565
7.3251 | -34.1882 | -13.1664
7.3312 | -34.1885 | -13.1766
7.3373 | -34.1889 | -13.1865
7.3434 | -34.1893 | -13.1965
7.3495 | -34.1897 | -13.2066
7.3556 | -34.19 | -13.2167
7.3617 | -34.1904 | -13.227
7.3678 | -34.1908 | -13.237
7.3739 | -34.1911 | -13.2474
7.3801 | -34.1915 | -13.2576
7.3862 | -34.1919 | -13.2677
7.3923 | -34.1923 | -13.278
7.3984 | -34.1926 | -13.2879
7.4045 | -34.193 | -13.298
7.4106 | -34.1934 | -13.308
7.4167 | -34.1938 | -13.3182
7.4228 | -34.1942 | -13.3285
7.4289 | -34.1945 | -13.3386
7.435 | -34.1949 | -13.3485
7.4411 | -34.1953 | -13.3587
7.4472 | -34.1956 | -13.3689
7.4533 | -34.196 | -13.3788
7.4594 | -34.1964 | -13.3892
7.4655 | -34.1968 | -13.3991
7.4716 | -34.1972 | -13.4094
7.4777 | -34.1975 | -13.4195
7.4838 | -34.1979 | -13.4295
7.4899 | -34.1983 | -13.4396
7.496 | -34.1986 | -13.4495
7.5021 | -34.199 | -13.4596
7.5082 | -34.1994 | -13.4696
7.5143 | -34.1998 | -13.4797
7.5204 | -34.2001 | -13.49
7.5266 | -34.2005 | -13.5
7.5327 | -34.2009 | -13.51
7.5388 | -34.2012 | -13.5204
7.5449 | -34.2017 | -13.5306
7.551 | -34.202 | -13.5406
7.5571 | -34.2024 | -13.5508
7.5632 | -34.2028 | -13.561
7.5693 | -34.2032 | -13.571
7.5754 | -34.2036 | -13.581
7.5815 | -34.2039 | -13.5911
7.5876 | -34.2043 | -13.6014
7.5937 | -34.2047 | -13.6117
7.5998 | -34.2051 | -13.6218
7.6059 | -34.2055 | -13.6321
7.612 | -34.2059 | -13.6422
7.6181 | -34.2063 | -13.6523
7.6242 | -34.2067 | -13.6627
7.6303 | -34.2071 | -13.6731
7.6364 | -34.2074 | -13.6833
7.6425 | -34.2078 | -13.6935
7.6486 | -34.2081 | -13.7036
7.6547 | -34.2085 | -13.7136
7.6608 | -34.2089 | -13.7237
7.667 | -34.2093 | -13.734
7.6731 | -34.2097 | -13.7442
7.6792 | -34.2101 | -13.7545
7.6853 | -34.2105 | -13.7647
7.6914 | -34.2108 | -13.775
7.6975 | -34.2112 | -13.7852
7.7036 | -34.2116 | -13.7953
7.7097 | -34.212 | -13.8057
7.7158 | -34.2124 | -13.8159
7.7219 | -34.2128 | -13.826
7.728 | -34.2132 | -13.8363
7.7341 | -34.2136 | -13.8465
7.7402 | -34.214 | -13.8568
7.7463 | -34.2144 | -13.8671
7.7524 | -34.2147 | -13.8776
7.7585 | -34.2151 | -13.8878
7.7646 | -34.2155 | -13.898
7.7707 | -34.2159 | -13.9084
7.7768 | -34.2162 | -13.9186
7.7829 | -34.2166 | -13.9291
7.789 | -34.217 | -13.9395
7.7951 | -34.2174 | -13.9497
7.8012 | -34.2178 | -13.9599
7.8073 | -34.2182 | -13.9703
7.8135 | -34.2186 | -13.9805
7.8196 | -34.219 | -13.9909
7.8257 | -34.2194 | -14.0014
7.8318 | -34.2198 | -14.0117
7.8379 | -34.2202 | -14.0218
7.844 | -34.2206 | -14.0321
7.8501 | -34.221 | -14.0424
7.8562 | -34.2214 | -14.0529
7.8623 | -34.2218 | -14.0631
7.8684 | -34.2222 | -14.0734
7.8745 | -34.2225 | -14.0836
7.8806 | -34.2229 | -14.094
7.8867 | -34.2233 | -14.1041
7.8928 | -34.2237 | -14.1141
7.8989 | -34.2241 | -14.1244
7.905 | -34.2245 | -14.1348
7.9111 | -34.2249 | -14.1451
7.9172 | -34.2253 | -14.1553
7.9233 | -34.2256 | -14.1655
7.9294 | -34.226 | -14.1759
7.9355 | -34.2264 | -14.1859
7.9416 | -34.2268 | -14.1962
7.9477 | -34.2272 | -14.2063
7.9539 | -34.2276 | -14.2166
7.96 | -34.228 | -14.2269
7.9661 | -34.2284 | -14.2373
7.9722 | -34.2288 | -14.2474
7.9783 | -34.2292 | -14.2579
7.9844 | -34.2296 | -14.2681
7.9905 | -34.23 | -14.2786
7.9966 | -34.2304 | -14.289
8.0027 | -34.2308 | -14.2992
8.0088 | -34.2312 | -14.3095
8.0149 | -34.2315 | -14.3198
8.021 | -34.2319 | -14.3301
8.0271 | -34.2323 | -14.3407
8.0332 | -34.2327 | -14.3509
8.0393 | -34.2331 | -14.3613
8.0454 | -34.2335 | -14.3719
8.0515 | -34.234 | -14.3821
8.0576 | -34.2343 | -14.3926
8.0637 | -34.2347 | -14.4028
8.0698 | -34.2351 | -14.413
8.0759 | -34.2355 | -14.4232
8.082 | -34.2359 | -14.4334
8.0881 | -34.2363 | -14.4438
8.0942 | -34.2367 | -14.4542
8.1004 | -34.2371 | -14.4646
8.1065 | -34.2375 | -14.4748
8.1126 | -34.2379 | -14.4851
8.1187 | -34.2383 | -14.4955
8.1248 | -34.2386 | -14.5056
8.1309 | -34.2391 | -14.5157
8.137 | -34.2394 | -14.5259
8.1431 | -34.2398 | -14.5364
8.1492 | -34.2402 | -14.5466
8.1553 | -34.2406 | -14.557
8.1614 | -34.241 | -14.5675
8.1675 | -34.2414 | -14.5777
8.1736 | -34.2417 | -14.5881
8.1797 | -34.2421 | -14.5985
8.1858 | -34.2425 | -14.6089
8.1919 | -34.2429 | -14.6192
8.198 | -34.2433 | -14.6293
8.2041 | -34.2438 | -14.6397
8.2102 | -34.2441 | -14.65
8.2163 | -34.2445 | -14.66
8.2224 | -34.2449 | -14.6703
8.2285 | -34.2453 | -14.6805
8.2346 | -34.2457 | -14.6907
8.2408 | -34.2461 | -14.7009
8.2469 | -34.2466 | -14.7112
8.253 | -34.247 | -14.7214
8.2591 | -34.2474 | -14.7318
8.2652 | -34.2478 | -14.742
8.2713 | -34.2482 | -14.7523
8.2774 | -34.2486 | -14.7627
8.2835 | -34.249 | -14.7733
8.2896 | -34.2494 | -14.7834
8.2957 | -34.2498 | -14.7937
8.3018 | -34.2502 | -14.8041
8.3079 | -34.2507 | -14.8141
8.314 | -34.2511 | -14.8244
8.3201 | -34.2515 | -14.8346
8.3262 | -34.2518 | -14.845
8.3323 | -34.2522 | -14.8551
8.3384 | -34.2527 | -14.8655
8.3445 | -34.253 | -14.8756
8.3506 | -34.2535 | -14.886
8.3567 | -34.2539 | -14.8964
8.3628 | -34.2543 | -14.9067
8.3689 | -34.2547 | -14.9168
8.375 | -34.2551 | -14.9271
8.3812 | -34.2555 | -14.9373
8.3873 | -34.2559 | -14.9475
8.3934 | -34.2563 | -14.9579
8.3995 | -34.2567 | -14.968
8.4056 | -34.2571 | -14.9782
8.4117 | -34.2574 | -14.9886
8.4178 | -34.2578 | -14.9989
8.4239 | -34.2582 | -15.009
8.43 | -34.2586 | -15.0193
8.4361 | -34.259 | -15.0296
8.4422 | -34.2594 | -15.0398
8.4483 | -34.2598 | -15.0499
8.4544 | -34.2602 | -15.0602
8.4605 | -34.2606 | -15.0702
8.4666 | -34.261 | -15.0807
8.4727 | -34.2614 | -15.091
8.4788 | -34.2618 | -15.1012
8.4849 | -34.2622 | -15.1115
8.491 | -34.2626 | -15.1217
8.4971 | -34.263 | -15.1317
8.5032 | -34.2634 | -15.1418
8.5093 | -34.2638 | -15.1518
8.5154 | -34.2642 | -15.1621
8.5215 | -34.2646 | -15.1721
8.5277 | -34.265 | -15.1823
8.5338 | -34.2654 | -15.1925
8.5399 | -34.2658 | -15.2027
8.546 | -34.2662 | -15.2127
8.5521 | -34.2666 | -15.2229
8.5582 | -34.267 | -15.2331
8.5643 | -34.2674 | -15.243
8.5704 | -34.2678 | -15.253
8.5765 | -34.2682 | -15.2633
8.5826 | -34.2686 | -15.2731
8.5887 | -34.269 | -15.2832
8.5948 | -34.2695 | -15.2934
8.6009 | -34.2699 | -15.3034
8.607 | -34.2703 | -15.3136
8.6131 | -34.2707 | -15.3238
8.6192 | -34.2711 | -15.3338
8.6253 | -34.2715 | -15.3439
8.6314 | -34.2719 | -15.354
8.6375 | -34.2723 | -15.3643
8.6436 | -34.2727 | -15.3744
8.6497 | -34.2731 | -15.3843
8.6558 | -34.2734 | -15.3944
8.6619 | -34.2738 | -15.4045
8.6681 | -34.2742 | -15.4148
8.6742 | -34.2746 | -15.4249
8.6803 | -34.275 | -15.435
8.6864 | -34.2755 | -15.4449
8.6925 | -34.2758 | -15.4549
8.6986 | -34.2762 | -15.465
8.7047 | -34.2766 | -15.4753
8.7108 | -34.277 | -15.4854
8.7169 | -34.2774 | -15.4955
8.723 | -34.2778 | -15.5054
8.7291 | -34.2782 | -15.5153
8.7352 | -34.2787 | -15.5255
8.7413 | -34.279 | -15.5357
8.7474 | -34.2794 | -15.5455
8.7535 | -34.2799 | -15.5555
8.7596 | -34.2803 | -15.5657
8.7657 | -34.2807 | -15.5758
8.7718 | -34.2811 | -15.5858
8.7779 | -34.2815 | -15.596
8.784 | -34.2819 | -15.6062
8.7901 | -34.2823 | -15.6162
8.7962 | -34.2827 | -15.626
8.8023 | -34.2831 | -15.6361
8.8084 | -34.2835 | -15.6461
8.8146 | -34.2839 | -15.6561
8.8207 | -34.2843 | -15.6661
8.8268 | -34.2848 | -15.6759
8.8329 | -34.2851 | -15.6859
8.839 | -34.2856 | -15.696
8.8451 | -34.286 | -15.706
8.8512 | -34.2864 | -15.7159
8.8573 | -34.2868 | -15.7259
8.8634 | -34.2872 | -15.7357
8.8695 | -34.2876 | -15.746
8.8756 | -34.288 | -15.756
8.8817 | -34.2884 | -15.7658
8.8878 | -34.2888 | -15.7761
8.8939 | -34.2892 | -15.7861
8.9 | -34.2896 | -15.796
8.9061 | -34.29 | -15.806
8.9122 | -34.2904 | -15.816
8.9183 | -34.2908 | -15.826
8.9244 | -34.2913 | -15.8362
8.9305 | -34.2917 | -15.846
8.9366 | -34.2921 | -15.856
8.9427 | -34.2925 | -15.8661
8.9488 | -34.2929 | -15.876
8.955 | -34.2933 | -15.886
8.9611 | -34.2938 | -15.8961
8.9672 | -34.2942 | -15.906
8.9733 | -34.2946 | -15.9158
8.9794 | -34.2951 | -15.9257
8.9855 | -34.2955 | -15.9358
8.9916 | -34.2959 | -15.9459
8.9977 | -34.2963 | -15.9559
9.0038 | -34.2967 | -15.9657
9.0099 | -34.2972 | -15.9758
9.016 | -34.2976 | -15.9858
9.0221 | -34.298 | -15.9956
9.0282 | -34.2984 | -16.0055
9.0343 | -34.2988 | -16.0156
9.0404 | -34.2993 | -16.0255
9.0465 | -34.2997 | -16.0356
9.0526 | -34.3001 | -16.0455
9.0587 | -34.3006 | -16.0554
9.0648 | -34.301 | -16.0652
9.0709 | -34.3014 | -16.0752
9.077 | -34.3019 | -16.085
9.0831 | -34.3023 | -16.0951
9.0892 | -34.3027 | -16.1052
9.0953 | -34.3031 | -16.1152
9.1015 | -34.3036 | -16.1251
9.1076 | -34.304 | -16.1348
9.1137 | -34.3044 | -16.1448
9.1198 | -34.3048 | -16.1547
9.1259 | -34.3052 | -16.1646
9.132 | -34.3057 | -16.1747
9.1381 | -34.3061 | -16.1845
9.1442 | -34.3066 | -16.1941
9.1503 | -34.307 | -16.2042
9.1564 | -34.3074 | -16.214
9.1625 | -34.3078 | -16.2239
9.1686 | -34.3083 | -16.2337
9.1747 | -34.3087 | -16.2436
9.1808 | -34.3092 | -16.2534
9.1869 | -34.3096 | -16.2633
9.193 | -34.31 | -16.2733
9.1991 | -34.3104 | -16.2832
9.2052 | -34.3109 | -16.293
9.2113 | -34.3113 | -16.3029
9.2174 | -34.3118 | -16.3127
9.2235 | -34.3122 | -16.3226
9.2296 | -34.3127 | -16.3325
9.2357 | -34.3131 | -16.3425
9.2419 | -34.3135 | -16.3524
9.248 | -34.314 | -16.3622
9.2541 | -34.3145 | -16.3722
9.2602 | -34.3149 | -16.3822
9.2663 | -34.3153 | -16.3917
9.2724 | -34.3158 | -16.4016
9.2785 | -34.3162 | -16.4115
9.2846 | -34.3167 | -16.4212
9.2907 | -34.3171 | -16.4311
9.2968 | -34.3175 | -16.441
9.3029 | -34.318 | -16.4511
9.309 | -34.3184 | -16.461
9.3151 | -34.3188 | -16.4708
9.3212 | -34.3193 | -16.4805
9.3273 | -34.3197 | -16.4904
9.3334 | -34.3202 | -16.5002
9.3395 | -34.3206 | -16.5098
9.3456 | -34.321 | -16.5194
9.3517 | -34.3214 | -16.5293
9.3578 | -34.3218 | -16.5391
9.3639 | -34.3223 | -16.5489
9.37 | -34.3228 | -16.5587
9.3761 | -34.3232 | -16.5687
9.3822 | -34.3236 | -16.5786
9.3884 | -34.3241 | -16.5881
9.3945 | -34.3245 | -16.598
9.4006 | -34.325 | -16.6078
9.4067 | -34.3254 | -16.6177
9.4128 | -34.3258 | -16.6278
9.4189 | -34.3263 | -16.6378
9.425 | -34.3267 | -16.6474
9.4311 | -34.3272 | -16.6573
9.4372 | -34.3276 | -16.667
9.4433 | -34.3281 | -16.6768
9.4494 | -34.3285 | -16.6864
9.4555 | -34.329 | -16.6962
9.4616 | -34.3294 | -16.706
9.4677 | -34.3299 | -16.7157
9.4738 | -34.3303 | -16.7254
9.4799 | -34.3308 | -16.7351
9.486 | -34.3312 | -16.745
9.4921 | -34.3317 | -16.7547
9.4982 | -34.3321 | -16.7644
9.5043 | -34.3326 | -16.7742
9.5104 | -34.333 | -16.7839
9.5165 | -34.3334 | -16.7937
9.5226 | -34.3339 | -16.8036
9.5288 | -34.3343 | -16.8132
9.5349 | -34.3348 | -16.8231
9.541 | -34.3352 | -16.8329
9.5471 | -34.3357 | -16.8427
9.5532 | -34.3361 | -16.8525
9.5593 | -34.3366 | -16.8625
9.5654 | -34.337 | -16.8722
9.5715 | -34.3375 | -16.8818
9.5776 | -34.3379 | -16.8917
9.5837 | -34.3384 | -16.9015
9.5898 | -34.3388 | -16.9115
9.5959 | -34.3392 | -16.9213
9.602 | -34.3397 | -16.9309
9.6081 | -34.3401 | -16.9406
9.6142 | -34.3406 | -16.9502
9.6203 | -34.341 | -16.96
9.6264 | -34.3414 | -16.9697
9.6325 | -34.3419 | -16.9795
9.6386 | -34.3423 | -16.9892
9.6447 | -34.3427 | -16.9989
9.6508 | -34.3432 | -17.0087
9.6569 | -34.3436 | -17.0185
9.663 | -34.3441 | -17.0284
9.6691 | -34.3445 | -17.0383
9.6753 | -34.345 | -17.0481
9.6814 | -34.3454 | -17.0578
9.6875 | -34.3459 | -17.0675
9.6936 | -34.3463 | -17.0772
9.6997 | -34.3468 | -17.0872
9.7058 | -34.3473 | -17.0967
9.7119 | -34.3477 | -17.1066
9.718 | -34.3482 | -17.1165
9.7241 | -34.3486 | -17.1261
9.7302 | -34.349 | -17.136
9.7363 | -34.3495 | -17.1457
9.7424 | -34.35 | -17.1553
9.7485 | -34.3504 | -17.165
9.7546 | -34.3508 | -17.1747
9.7607 | -34.3513 | -17.1846
9.7668 | -34.3517 | -17.1943
9.7729 | -34.3522 | -17.2041
9.779 | -34.3527 | -17.2136
9.7851 | -34.3531 | -17.2234
9.7912 | -34.3536 | -17.2331
9.7973 | -34.3541 | -17.243
9.8034 | -34.3546 | -17.2526
9.8095 | -34.355 | -17.2625
9.8157 | -34.3555 | -17.2724
9.8218 | -34.3559 | -17.2824
9.8279 | -34.3564 | -17.292
9.834 | -34.3568 | -17.3019
9.8401 | -34.3573 | -17.3117
9.8462 | -34.3577 | -17.3213
9.8523 | -34.3582 | -17.331
9.8584 | -34.3586 | -17.3407
9.8645 | -34.3591 | -17.3505
9.8706 | -34.3595 | -17.3604
9.8767 | -34.3599 | -17.3702
9.8828 | -34.3603 | -17.3798
9.8889 | -34.3608 | -17.3895
9.895 | -34.3612 | -17.3992
9.9011 | -34.3617 | -17.409
9.9072 | -34.3621 | -17.4188
9.9133 | -34.3625 | -17.4285
9.9194 | -34.363 | -17.4381
9.9255 | -34.3635 | -17.4478
9.9316 | -34.3639 | -17.4575
9.9377 | -34.3644 | -17.4671
9.9438 | -34.3648 | -17.4769
9.9499 | -34.3652 | -17.4868
9.956 | -34.3657 | -17.4965
9.9622 | -34.3661 | -17.5063
9.9683 | -34.3666 | -17.5161
9.9744 | -34.367 | -17.5258
9.9805 | -34.3675 | -17.5355
9.9866 | -34.3679 | -17.5452
9.9927 | -34.3683 | -17.5548
9.9988 | -34.3688 | -17.5647
10.0049 | -34.3692 | -17.5743
10.011 | -34.3696 | -17.5842
10.0171 | -34.3701 | -17.5937
10.0232 | -34.3706 | -17.6033
10.0293 | -34.371 | -17.6131
10.0354 | -34.3715 | -17.6229
10.0415 | -34.3719 | -17.6327
10.0476 | -34.3724 | -17.6425
10.0537 | -34.3728 | -17.6523
10.0598 | -34.3733 | -17.662
10.0659 | -34.3737 | -17.6718
10.072 | -34.3741 | -17.6813
10.0781 | -34.3746 | -17.691
10.0842 | -34.375 | -17.7007
10.0903 | -34.3755 | -17.7103
10.0964 | -34.3759 | -17.7199
10.1026 | -34.3763 | -17.7297
10.1087 | -34.3768 | -17.7395
10.1148 | -34.3772 | -17.7492
10.1209 | -34.3777 | -17.7589
10.127 | -34.3781 | -17.7686
10.1331 | -34.3786 | -17.7782
10.1392 | -34.379 | -17.7876
10.1453 | -34.3794 | -17.7972
10.1514 | -34.3799 | -17.807
10.1575 | -34.3803 | -17.8168
10.1636 | -34.3808 | -17.8263
10.1697 | -34.3812 | -17.836
10.1758 | -34.3817 | -17.8455
10.1819 | -34.3821 | -17.8551
10.188 | -34.3825 | -17.8647
10.1941 | -34.383 | -17.8745
10.2002 | -34.3834 | -17.8842
10.2063 | -34.3839 | -17.8938
10.2124 | -34.3843 | -17.9037
10.2185 | -34.3847 | -17.9135
10.2246 | -34.3852 | -17.9232
10.2307 | -34.3857 | -17.9328
10.2368 | -34.3861 | -17.9422
10.2429 | -34.3865 | -17.9518
10.2491 | -34.3869 | -17.9616
10.2552 | -34.3874 | -17.9713
10.2613 | -34.3878 | -17.981
10.2674 | -34.3882 | -17.9908
10.2735 | -34.3887 | -18.0007
10.2796 | -34.3891 | -18.0104
10.2857 | -34.3896 | -18.0202
10.2918 | -34.39 | -18.03
10.2979 | -34.3904 | -18.0398
10.304 | -34.3909 | -18.0493
10.3101 | -34.3913 | -18.059
10.3162 | -34.3918 | -18.0688
10.3223 | -34.3922 | -18.0783
10.3284 | -34.3926 | -18.088
10.3345 | -34.3931 | -18.0976
10.3406 | -34.3935 | -18.1073
10.3467 | -34.3939 | -18.1169
10.3528 | -34.3943 | -18.1265
10.3589 | -34.3948 | -18.1363
10.365 | -34.3952 | -18.1461
10.3711 | -34.3957 | -18.1559
10.3772 | -34.3961 | -18.1657
10.3833 | -34.3966 | -18.1754
10.3895 | -34.3971 | -18.1851
10.3956 | -34.3975 | -18.1949
10.4017 | -34.3979 | -18.2045
10.4078 | -34.3984 | -18.2142
10.4139 | -34.3988 | -18.2239
10.42 | -34.3993 | -18.2337
10.4261 | -34.3997 | -18.2433
10.4322 | -34.4001 | -18.2529
10.4383 | -34.4005 | -18.2625
10.4444 | -34.401 | -18.2721
10.4505 | -34.4014 | -18.2819
10.4566 | -34.4018 | -18.2914
10.4627 | -34.4023 | -18.3009
10.4688 | -34.4027 | -18.3105
10.4749 | -34.4032 | -18.3203
10.481 | -34.4037 | -18.3299
10.4871 | -34.4041 | -18.3394
10.4932 | -34.4045 | -18.349
10.4993 | -34.405 | -18.3587
10.5054 | -34.4054 | -18.3685
10.5115 | -34.4059 | -18.3781
10.5176 | -34.4063 | -18.3875
10.5237 | -34.4068 | -18.3972
10.5298 | -34.4072 | -18.4067
10.536 | -34.4076 | -18.4165
10.5421 | -34.4081 | -18.4261
10.5482 | -34.4085 | -18.4358
10.5543 | -34.409 | -18.4455
10.5604 | -34.4094 | -18.455
10.5665 | -34.4099 | -18.4645
10.5726 | -34.4103 | -18.4741
10.5787 | -34.4107 | -18.4838
10.5848 | -34.4112 | -18.4929
10.5909 | -34.4116 | -18.5027
10.597 | -34.4121 | -18.5125
10.6031 | -34.4125 | -18.5224
10.6092 | -34.4129 | -18.532
10.6153 | -34.4134 | -18.5417
10.6214 | -34.4138 | -18.5513
10.6275 | -34.4143 | -18.561
10.6336 | -34.4147 | -18.5705
10.6397 | -34.4151 | -18.5803
10.6458 | -34.4156 | -18.5898
10.6519 | -34.416 | -18.5996
10.658 | -34.4165 | -18.6093
10.6641 | -34.4169 | -18.6191
10.6702 | -34.4173 | -18.6286
10.6764 | -34.4178 | -18.6383
10.6825 | -34.4183 | -18.6479
10.6886 | -34.4187 | -18.6573
10.6947 | -34.4192 | -18.6668
10.7008 | -34.4196 | -18.6766
10.7069 | -34.42 | -18.6864
10.713 | -34.4205 | -18.6959
10.7191 | -34.4209 | -18.7054
10.7252 | -34.4213 | -18.7148
10.7313 | -34.4218 | -18.7245
10.7374 | -34.4222 | -18.734
10.7435 | -34.4226 | -18.7435
10.7496 | -34.4231 | -18.753
10.7557 | -34.4235 | -18.7627
10.7618 | -34.4239 | -18.7721
10.7679 | -34.4244 | -18.7817
10.774 | -34.4248 | -18.7913
10.7801 | -34.4253 | -18.801
10.7862 | -34.4257 | -18.8105
10.7923 | -34.4261 | -18.8202
10.7984 | -34.4266 | -18.8299
10.8045 | -34.427 | -18.8392
10.8106 | -34.4275 | -18.8488
10.8168 | -34.4279 | -18.8584
10.8229 | -34.4283 | -18.8682
10.829 | -34.4288 | -18.8778
10.8351 | -34.4292 | -18.887
10.8412 | -34.4297 | -18.8966
10.8473 | -34.4301 | -18.9061
10.8534 | -34.4306 | -18.9159
10.8595 | -34.4311 | -18.9254
10.8656 | -34.4316 | -18.9347
10.8717 | -34.432 | -18.9443
10.8778 | -34.4324 | -18.9539
10.8839 | -34.4329 | -18.9635
10.89 | -34.4333 | -18.9733
10.8961 | -34.4338 | -18.9827
10.9022 | -34.4342 | -18.992
10.9083 | -34.4347 | -19.0014
10.9144 | -34.4351 | -19.011
10.9205 | -34.4356 | -19.0204
10.9266 | -34.436 | -19.0298
10.9327 | -34.4365 | -19.0393
10.9388 | -34.437 | -19.0485
10.9449 | -34.4374 | -19.058
10.951 | -34.4378 | -19.0674
10.9571 | -34.4383 | -19.0769
10.9633 | -34.4387 | -19.0866
10.9694 | -34.4392 | -19.0962
10.9755 | -34.4397 | -19.1057
10.9816 | -34.4401 | -19.1153
10.9877 | -34.4406 | -19.1247
10.9938 | -34.441 | -19.1344
10.9999 | -34.4415 | -19.1438
11.006 | -34.442 | -19.1532
11.0121 | -34.4424 | -19.1627
11.0182 | -34.4428 | -19.1722
11.0243 | -34.4433 | -19.1816
11.0304 | -34.4437 | -19.1911
11.0365 | -34.4442 | -19.2006
11.0426 | -34.4446 | -19.2102
11.0487 | -34.4451 | -19.2198
11.0548 | -34.4456 | -19.2292
11.0609 | -34.446 | -19.239
11.067 | -34.4465 | -19.2485
11.0731 | -34.447 | -19.2577
11.0792 | -34.4474 | -19.2671
11.0853 | -34.4479 | -19.2767
11.0914 | -34.4484 | -19.2863
11.0975 | -34.4488 | -19.2959
11.1037 | -34.4492 | -19.3052
11.1098 | -34.4497 | -19.3149
11.1159 | -34.4501 | -19.3242
11.122 | -34.4505 | -19.3335
11.1281 | -34.451 | -19.343
11.1342 | -34.4514 | -19.3526
11.1403 | -34.4519 | -19.3621
11.1464 | -34.4524 | -19.3717
11.1525 | -34.4528 | -19.3809
11.1586 | -34.4532 | -19.3904
11.1647 | -34.4537 | -19.3997
11.1708 | -34.4542 | -19.409
11.1769 | -34.4546 | -19.4184
11.183 | -34.4551 | -19.4279
11.1891 | -34.4556 | -19.4374
11.1952 | -34.456 | -19.4469
11.2013 | -34.4565 | -19.4563
11.2074 | -34.457 | -19.4658
11.2135 | -34.4574 | -19.4755
11.2196 | -34.4579 | -19.485
11.2257 | -34.4584 | -19.4946
11.2318 | -34.4588 | -19.504
11.2379 | -34.4593 | -19.5135
11.244 | -34.4598 | -19.5232
11.2502 | -34.4603 | -19.5325
11.2563 | -34.4608 | -19.5418
11.2624 | -34.4612 | -19.5512
11.2685 | -34.4617 | -19.5607
11.2746 | -34.4622 | -19.57
11.2807 | -34.4626 | -19.5797
11.2868 | -34.4631 | -19.5893
11.2929 | -34.4636 | -19.5987
11.299 | -34.464 | -19.6081
11.3051 | -34.4645 | -19.6177
11.3112 | -34.465 | -19.6272
11.3173 | -34.4655 | -19.6364
11.3234 | -34.466 | -19.6458
11.3295 | -34.4664 | -19.6553
11.3356 | -34.4668 | -19.6649
11.3417 | -34.4673 | -19.6741
11.3478 | -34.4677 | -19.6837
11.3539 | -34.4682 | -19.6932
11.36 | -34.4687 | -19.7025
11.3661 | -34.4692 | -19.7117
11.3722 | -34.4696 | -19.7209
11.3783 | -34.4701 | -19.7303
11.3844 | -34.4706 | -19.7399
11.3906 | -34.471 | -19.7494
11.3967 | -34.4715 | -19.7585
11.4028 | -34.472 | -19.7681
11.4089 | -34.4724 | -19.7773
11.415 | -34.4729 | -19.7866
11.4211 | -34.4733 | -19.7961
11.4272 | -34.4738 | -19.8056
11.4333 | -34.4743 | -19.8144
11.4394 | -34.4748 | -19.8238
11.4455 | -34.4752 | -19.8333
11.4516 | -34.4757 | -19.8426
11.4577 | -34.4762 | -19.8521
11.4638 | -34.4767 | -19.8613
11.4699 | -34.4772 | -19.8708
11.476 | -34.4776 | -19.8804
11.4821 | -34.4781 | -19.8896
11.4882 | -34.4787 | -19.8993
11.4943 | -34.4791 | -19.9087
11.5004 | -34.4796 | -19.9178
11.5065 | -34.4801 | -19.927
11.5126 | -34.4806 | -19.9364
11.5187 | -34.4811 | -19.9458
11.5248 | -34.4815 | -19.9553
11.5309 | -34.482 | -19.9646
11.5371 | -34.4825 | -19.9739
11.5432 | -34.483 | -19.9834
11.5493 | -34.4835 | -19.9929
11.5554 | -34.484 | -20.0022
11.5615 | -34.4845 | -20.0117
11.5676 | -34.485 | -20.0211
11.5737 | -34.4854 | -20.0304
11.5798 | -34.4859 | -20.0395
11.5859 | -34.4864 | -20.0489
11.592 | -34.4869 | -20.058
11.5981 | -34.4874 | -20.0672
11.6042 | -34.4879 | -20.0766
11.6103 | -34.4884 | -20.0858
11.6164 | -34.4889 | -20.0951
11.6225 | -34.4893 | -20.1044
11.6286 | -34.4898 | -20.1136
11.6347 | -34.4903 | -20.123
11.6408 | -34.4908 | -20.1323
11.6469 | -34.4913 | -20.1416
11.653 | -34.4918 | -20.1507
11.6591 | -34.4923 | -20.1601
11.6652 | -34.4927 | -20.1693
11.6713 | -34.4933 | -20.1786
11.6775 | -34.4938 | -20.1877
11.6836 | -34.4943 | -20.197
11.6897 | -34.4948 | -20.2063
11.6958 | -34.4953 | -20.2156
11.7019 | -34.4958 | -20.225
11.708 | -34.4962 | -20.2342
11.7141 | -34.4968 | -20.2433
11.7202 | -34.4973 | -20.2526
11.7263 | -34.4977 | -20.2618
11.7324 | -34.4982 | -20.2713
11.7385 | -34.4987 | -20.2806
11.7446 | -34.4991 | -20.2898
11.7507 | -34.4997 | -20.2989
11.7568 | -34.5001 | -20.3082
11.7629 | -34.5006 | -20.3176
11.769 | -34.5011 | -20.3268
11.7751 | -34.5016 | -20.336
11.7812 | -34.5021 | -20.3451
11.7873 | -34.5026 | -20.3544
11.7934 | -34.5031 | -20.3638
11.7995 | -34.5036 | -20.373
11.8056 | -34.5041 | -20.3821
11.8117 | -34.5046 | -20.3914
11.8178 | -34.5051 | -20.4004
11.824 | -34.5056 | -20.4095
11.8301 | -34.5061 | -20.4187
11.8362 | -34.5066 | -20.4278
11.8423 | -34.5071 | -20.4371
11.8484 | -34.5076 | -20.4464
11.8545 | -34.5081 | -20.4555
11.8606 | -34.5086 | -20.4646
11.8667 | -34.5091 | -20.474
11.8728 | -34.5096 | -20.483
11.8789 | -34.5101 | -20.4922
11.885 | -34.5106 | -20.5015
11.8911 | -34.5111 | -20.5105
11.8972 | -34.5116 | -20.5198
11.9033 | -34.5121 | -20.5289
11.9094 | -34.5126 | -20.538
11.9155 | -34.5131 | -20.5472
11.9216 | -34.5136 | -20.5561
11.9277 | -34.5141 | -20.5653
11.9338 | -34.5146 | -20.5745
11.9399 | -34.5151 | -20.5838
11.946 | -34.5156 | -20.5929
11.9521 | -34.5161 | -20.602
11.9582 | -34.5166 | -20.6112
11.9644 | -34.5171 | -20.6202
11.9705 | -34.5177 | -20.6294
11.9766 | -34.5182 | -20.6385
11.9827 | -34.5187 | -20.6476
11.9888 | -34.5192 | -20.6567
11.9949 | -34.5198 | -20.6655
12.001 | -34.5203 | -20.6748
12.0071 | -34.5208 | -20.6839
12.0132 | -34.5213 | -20.6929
12.0193 | -34.5217 | -20.7021
12.0254 | -34.5222 | -20.7113
12.0315 | -34.5228 | -20.7202
12.0376 | -34.5233 | -20.7293
12.0437 | -34.5238 | -20.7384
12.0498 | -34.5243 | -20.7476
12.0559 | -34.5248 | -20.7566
12.062 | -34.5253 | -20.7656
12.0681 | -34.5258 | -20.7748
12.0742 | -34.5263 | -20.7838
12.0803 | -34.5268 | -20.7931
12.0864 | -34.5273 | -20.8022
12.0925 | -34.5278 | -20.8113
12.0986 | -34.5284 | -20.8204
12.1047 | -34.5289 | -20.8296
12.1109 | -34.5294 | -20.8386
12.117 | -34.5299 | -20.8478
12.1231 | -34.5305 | -20.8569
12.1292 | -34.531 | -20.8659
12.1353 | -34.5315 | -20.8749
12.1414 | -34.532 | -20.8841
12.1475 | -34.5325 | -20.8933
12.1536 | -34.533 | -20.9024
12.1597 | -34.5335 | -20.9115
12.1658 | -34.5341 | -20.9205
12.1719 | -34.5346 | -20.9294
12.178 | -34.5351 | -20.9385
12.1841 | -34.5356 | -20.9477
12.1902 | -34.5361 | -20.9569
12.1963 | -34.5366 | -20.9661
12.2024 | -34.5372 | -20.9749
12.2085 | -34.5377 | -20.9841
12.2146 | -34.5382 | -20.9932
12.2207 | -34.5387 | -21.0022
12.2268 | -34.5393 | -21.0113
12.2329 | -34.5398 | -21.0204
12.239 | -34.5403 | -21.029
12.2451 | -34.5408 | -21.0381
12.2513 | -34.5413 | -21.0471
12.2574 | -34.5419 | -21.0561
12.2635 | -34.5424 | -21.0651
12.2696 | -34.5429 | -21.0741
12.2757 | -34.5434 | -21.083
12.2818 | -34.544 | -21.092
12.2879 | -34.5445 | -21.1011
12.294 | -34.545 | -21.11
12.3001 | -34.5455 | -21.1191
12.3062 | -34.5461 | -21.1281
12.3123 | -34.5466 | -21.1369
12.3184 | -34.5471 | -21.1458
12.3245 | -34.5476 | -21.1546
12.3306 | -34.5481 | -21.1635
12.3367 | -34.5487 | -21.1724
12.3428 | -34.5492 | -21.1814
12.3489 | -34.5497 | -21.1902
12.355 | -34.5502 | -21.1993
12.3611 | -34.5508 | -21.2084
12.3672 | -34.5512 | -21.2173
12.3733 | -34.5518 | -21.2263
12.3794 | -34.5524 | -21.235
12.3855 | -34.5529 | -21.2438
12.3916 | -34.5534 | -21.2528
12.3978 | -34.5539 | -21.2619
12.4039 | -34.5544 | -21.271
12.41 | -34.555 | -21.2801
12.4161 | -34.5555 | -21.2887
12.4222 | -34.556 | -21.2976
12.4283 | -34.5566 | -21.3065
12.4344 | -34.5571 | -21.3156
12.4405 | -34.5576 | -21.3247
12.4466 | -34.5581 | -21.3335
12.4527 | -34.5586 | -21.3426
12.4588 | -34.5591 | -21.3515
12.4649 | -34.5596 | -21.3604
12.471 | -34.5602 | -21.3695
12.4771 | -34.5607 | -21.3783
12.4832 | -34.5612 | -21.3873
12.4893 | -34.5617 | -21.3961
12.4954 | -34.5623 | -21.4048
12.5015 | -34.5628 | -21.4137
12.5076 | -34.5633 | -21.4227
12.5137 | -34.5638 | -21.4317
12.5198 | -34.5644 | -21.4405
12.5259 | -34.5649 | -21.4495
12.532 | -34.5654 | -21.4583
12.5382 | -34.5659 | -21.4673
12.5443 | -34.5664 | -21.4762
12.5504 | -34.5669 | -21.4851
12.5565 | -34.5675 | -21.4938
12.5626 | -34.568 | -21.5029
12.5687 | -34.5685 | -21.5118
12.5748 | -34.5691 | -21.5206
12.5809 | -34.5696 | -21.5295
12.587 | -34.5701 | -21.5386
12.5931 | -34.5706 | -21.5472
12.5992 | -34.5712 | -21.556
12.6053 | -34.5717 | -21.5649
12.6114 | -34.5722 | -21.5737
12.6175 | -34.5728 | -21.5827
12.6236 | -34.5733 | -21.5918
12.6297 | -34.5738 | -21.6005
12.6358 | -34.5743 | -21.6093
12.6419 | -34.5748 | -21.6182
12.648 | -34.5754 | -21.6272
12.6541 | -34.5759 | -21.6361
12.6602 | -34.5764 | -21.6445
12.6663 | -34.5769 | -21.6534
12.6724 | -34.5775 | -21.6622
12.6785 | -34.578 | -21.671
12.6847 | -34.5785 | -21.6798
12.6908 | -34.5791 | -21.6887
12.6969 | -34.5796 | -21.6976
12.703 | -34.5801 | -21.7065
12.7091 | -34.5806 | -21.7154
12.7152 | -34.5811 | -21.7241
12.7213 | -34.5816 | -21.733
12.7274 | -34.5822 | -21.7419
12.7335 | -34.5827 | -21.7503
12.7396 | -34.5832 | -21.759
12.7457 | -34.5837 | -21.768
12.7518 | -34.5842 | -21.7767
12.7579 | -34.5848 | -21.7857
12.764 | -34.5853 | -21.7946
12.7701 | -34.5858 | -21.8033
12.7762 | -34.5863 | -21.8119
12.7823 | -34.5869 | -21.8208
12.7884 | -34.5874 | -21.8296
12.7945 | -34.5879 | -21.8385
12.8006 | -34.5885 | -21.8471
12.8067 | -34.589 | -21.8559
12.8128 | -34.5895 | -21.8645
12.8189 | -34.59 | -21.8732
12.8251 | -34.5905 | -21.882
12.8312 | -34.591 | -21.8909
12.8373 | -34.5915 | -21.8994
12.8434 | -34.592 | -21.9083
12.8495 | -34.5926 | -21.9169
12.8556 | -34.5931 | -21.9255
12.8617 | -34.5936 | -21.9344
12.8678 | -34.5941 | -21.9434
12.8739 | -34.5946 | -21.9522
12.88 | -34.5951 | -21.9609
12.8861 | -34.5956 | -21.9696
12.8922 | -34.5962 | -21.9785
12.8983 | -34.5967 | -21.9874
12.9044 | -34.5972 | -21.9959
12.9105 | -34.5977 | -22.0045
12.9166 | -34.5982 | -22.0133
12.9227 | -34.5987 | -22.022
12.9288 | -34.5993 | -22.0309
12.9349 | -34.5998 | -22.0398
12.941 | -34.6003 | -22.0485
12.9471 | -34.6008 | -22.0573
12.9532 | -34.6014 | -22.0661
12.9593 | -34.6019 | -22.0751
12.9654 | -34.6024 | -22.084
12.9716 | -34.6029 | -22.0931
12.9777 | -34.6035 | -22.1019
12.9838 | -34.604 | -22.1105
12.9899 | -34.6045 | -22.1193
12.996 | -34.605 | -22.128
13.0021 | -34.6056 | -22.1368
13.0082 | -34.6061 | -22.1454
13.0143 | -34.6066 | -22.1538
13.0204 | -34.6071 | -22.1623
13.0265 | -34.6077 | -22.1709
13.0326 | -34.6082 | -22.1796
13.0387 | -34.6087 | -22.1885
13.0448 | -34.6092 | -22.1972
13.0509 | -34.6097 | -22.2055
13.057 | -34.6102 | -22.2142
13.0631 | -34.6108 | -22.2229
13.0692 | -34.6113 | -22.2315
13.0753 | -34.6118 | -22.2404
13.0814 | -34.6123 | -22.2495
13.0875 | -34.6129 | -22.258
13.0936 | -34.6134 | -22.267
13.0997 | -34.6139 | -22.2756
13.1058 | -34.6144 | -22.2842
13.112 | -34.6149 | -22.2927
13.1181 | -34.6154 | -22.3013
13.1242 | -34.616 | -22.3101
13.1303 | -34.6165 | -22.3187
13.1364 | -34.617 | -22.3276
13.1425 | -34.6175 | -22.3362
13.1486 | -34.618 | -22.345
13.1547 | -34.6185 | -22.3537
13.1608 | -34.619 | -22.3623
13.1669 | -34.6196 | -22.3709
13.173 | -34.6201 | -22.3798
13.1791 | -34.6206 | -22.3885
13.1852 | -34.6211 | -22.3971
13.1913 | -34.6216 | -22.4057
13.1974 | -34.6221 | -22.4144
13.2035 | -34.6226 | -22.4229
13.2096 | -34.6232 | -22.4315
13.2157 | -34.6237 | -22.44
13.2218 | -34.6242 | -22.4486
13.2279 | -34.6247 | -22.4571
13.234 | -34.6253 | -22.4656
13.2401 | -34.6258 | -22.4741
13.2462 | -34.6263 | -22.4827
13.2524 | -34.6268 | -22.4915
13.2585 | -34.6273 | -22.5
13.2646 | -34.6278 | -22.5084
13.2707 | -34.6284 | -22.5172
13.2768 | -34.6289 | -22.5261
13.2829 | -34.6294 | -22.5349
13.289 | -34.6299 | -22.5437
13.2951 | -34.6304 | -22.5522
13.3012 | -34.6309 | -22.561
13.3073 | -34.6315 | -22.5696
13.3134 | -34.632 | -22.5783
13.3195 | -34.6325 | -22.5868
13.3256 | -34.633 | -22.5955
13.3317 | -34.6335 | -22.604
13.3378 | -34.634 | -22.6128
13.3439 | -34.6345 | -22.6216
13.35 | -34.635 | -22.6303
13.3561 | -34.6355 | -22.6388
13.3622 | -34.6361 | -22.6474
13.3683 | -34.6366 | -22.6558
13.3744 | -34.6371 | -22.6644
13.3805 | -34.6376 | -22.673
13.3866 | -34.6381 | -22.6815
13.3927 | -34.6386 | -22.6903
13.3989 | -34.6391 | -22.6988
13.405 | -34.6396 | -22.7074
13.4111 | -34.6401 | -22.7159
13.4172 | -34.6406 | -22.7245
13.4233 | -34.6411 | -22.7333
13.4294 | -34.6416 | -22.7422
13.4355 | -34.6421 | -22.7505
13.4416 | -34.6426 | -22.7591
13.4477 | -34.6432 | -22.7676
13.4538 | -34.6437 | -22.7764
13.4599 | -34.6442 | -22.785
13.466 | -34.6447 | -22.7938
13.4721 | -34.6452 | -22.8021
13.4782 | -34.6457 | -22.8106
13.4843 | -34.6463 | -22.8195
13.4904 | -34.6468 | -22.8279
13.4965 | -34.6473 | -22.8366
13.5026 | -34.6478 | -22.8453
13.5087 | -34.6483 | -22.8537
13.5148 | -34.6488 | -22.8624
13.5209 | -34.6494 | -22.8711
13.527 | -34.6499 | -22.8796
13.5331 | -34.6504 | -22.888
13.5393 | -34.6509 | -22.8964
13.5454 | -34.6514 | -22.905
13.5515 | -34.6519 | -22.9139
13.5576 | -34.6524 | -22.9225
13.5637 | -34.653 | -22.9311
13.5698 | -34.6535 | -22.9396
13.5759 | -34.654 | -22.9479
13.582 | -34.6545 | -22.9565
13.5881 | -34.655 | -22.9652
13.5942 | -34.6555 | -22.974
13.6003 | -34.656 | -22.9826
13.6064 | -34.6565 | -22.9915
13.6125 | -34.657 | -22.9999
13.6186 | -34.6575 | -23.0084
13.6247 | -34.6581 | -23.0171
13.6308 | -34.6585 | -23.0257
13.6369 | -34.659 | -23.0342
13.643 | -34.6596 | -23.0428
13.6491 | -34.6601 | -23.0514
13.6552 | -34.6606 | -23.0599
13.6613 | -34.6611 | -23.0686
13.6674 | -34.6616 | -23.0772
13.6735 | -34.662 | -23.0857
13.6796 | -34.6626 | -23.0941
13.6858 | -34.663 | -23.1027
13.6919 | -34.6636 | -23.1113
13.698 | -34.664 | -23.1198
13.7041 | -34.6645 | -23.1285
13.7102 | -34.6651 | -23.1373
13.7163 | -34.6656 | -23.1458
13.7224 | -34.6661 | -23.1543
13.7285 | -34.6666 | -23.1629
13.7346 | -34.6671 | -23.1714
13.7407 | -34.6677 | -23.1798
13.7468 | -34.6682 | -23.1883
13.7529 | -34.6687 | -23.1968
13.759 | -34.6692 | -23.2056
13.7651 | -34.6697 | -23.214
13.7712 | -34.6702 | -23.2224
13.7773 | -34.6707 | -23.2309
13.7834 | -34.6712 | -23.2395
13.7895 | -34.6717 | -23.2481
13.7956 | -34.6722 | -23.2567
13.8017 | -34.6727 | -23.2653
13.8078 | -34.6732 | -23.2738
13.8139 | -34.6737 | -23.2824
13.82 | -34.6742 | -23.2909
13.8262 | -34.6747 | -23.2995
13.8323 | -34.6752 | -23.3081
13.8384 | -34.6757 | -23.3166
13.8445 | -34.6762 | -23.325
13.8506 | -34.6768 | -23.3336
13.8567 | -34.6773 | -23.3422
13.8628 | -34.6778 | -23.3508
13.8689 | -34.6783 | -23.3593
13.875 | -34.6788 | -23.3679
13.8811 | -34.6793 | -23.3765
13.8872 | -34.6799 | -23.3851
13.8933 | -34.6804 | -23.3935
13.8994 | -34.6809 | -23.4022
13.9055 | -34.6813 | -23.411
13.9116 | -34.6818 | -23.4196
13.9177 | -34.6823 | -23.4281
13.9238 | -34.6828 | -23.4367
13.9299 | -34.6833 | -23.4452
13.936 | -34.6838 | -23.4537
13.9421 | -34.6843 | -23.4622
13.9482 | -34.6848 | -23.4706
13.9543 | -34.6853 | -23.479
13.9604 | -34.6858 | -23.4876
13.9665 | -34.6863 | -23.496
13.9727 | -34.6868 | -23.5046
13.9788 | -34.6874 | -23.513
13.9849 | -34.6879 | -23.5214
13.991 | -34.6884 | -23.5299
13.9971 | -34.6889 | -23.5381
14.0032 | -34.6894 | -23.5466
14.0093 | -34.6899 | -23.5552
14.0154 | -34.6904 | -23.5636
14.0215 | -34.6909 | -23.572
14.0276 | -34.6914 | -23.5805
14.0337 | -34.6919 | -23.589
14.0398 | -34.6925 | -23.5975
14.0459 | -34.693 | -23.606
14.052 | -34.6935 | -23.6146
14.0581 | -34.694 | -23.6233
14.0642 | -34.6945 | -23.6318
14.0703 | -34.695 | -23.6405
14.0764 | -34.6955 | -23.649
14.0825 | -34.696 | -23.6577
14.0886 | -34.6966 | -23.6662
14.0947 | -34.6971 | -23.6747
14.1008 | -34.6976 | -23.6833
14.1069 | -34.6981 | -23.6916
14.1131 | -34.6985 | -23.7001
14.1192 | -34.6991 | -23.7086
14.1253 | -34.6996 | -23.717
14.1314 | -34.7001 | -23.7256
14.1375 | -34.7006 | -23.7339
14.1436 | -34.7012 | -23.7425
14.1497 | -34.7017 | -23.7511
14.1558 | -34.7022 | -23.7596
14.1619 | -34.7027 | -23.768
14.168 | -34.7032 | -23.7763
14.1741 | -34.7037 | -23.7849
14.1802 | -34.7042 | -23.7934
14.1863 | -34.7047 | -23.802
14.1924 | -34.7053 | -23.8107
14.1985 | -34.7058 | -23.8192
14.2046 | -34.7063 | -23.8277
14.2107 | -34.7068 | -23.836
14.2168 | -34.7073 | -23.8446
14.2229 | -34.7078 | -23.8532
14.229 | -34.7083 | -23.8618
14.2351 | -34.7088 | -23.8703
14.2412 | -34.7093 | -23.8789
14.2473 | -34.7099 | -23.8873
14.2534 | -34.7104 | -23.8957
14.2596 | -34.7109 | -23.9041
14.2657 | -34.7114 | -23.9125
14.2718 | -34.7119 | -23.9211
14.2779 | -34.7124 | -23.9295
14.284 | -34.713 | -23.9378
14.2901 | -34.7135 | -23.9464
14.2962 | -34.714 | -23.9549
14.3023 | -34.7145 | -23.9635
14.3084 | -34.715 | -23.9721
14.3145 | -34.7155 | -23.9805
14.3206 | -34.716 | -23.9892
14.3267 | -34.7165 | -23.9977
14.3328 | -34.717 | -24.0063
14.3389 | -34.7175 | -24.0147
14.345 | -34.718 | -24.0231
14.3511 | -34.7185 | -24.0317
14.3572 | -34.7191 | -24.0403
14.3633 | -34.7196 | -24.049
14.3694 | -34.7201 | -24.0573
14.3755 | -34.7206 | -24.0658
14.3816 | -34.7212 | -24.0745
14.3877 | -34.7217 | -24.0828
14.3938 | -34.7222 | -24.0915
14.4 | -34.7228 | -24.0998
14.4061 | -34.7233 | -24.1083
14.4122 | -34.7238 | -24.1169
14.4183 | -34.7243 | -24.1253
14.4244 | -34.7248 | -24.1336
14.4305 | -34.7253 | -24.1421
14.4366 | -34.7258 | -24.1506
14.4427 | -34.7264 | -24.1593
14.4488 | -34.7269 | -24.1678
14.4549 | -34.7274 | -24.1762
14.461 | -34.7279 | -24.1846
14.4671 | -34.7284 | -24.1932
14.4732 | -34.729 | -24.2019
14.4793 | -34.7295 | -24.2104
14.4854 | -34.73 | -24.2188
14.4915 | -34.7305 | -24.2271
14.4976 | -34.7311 | -24.2354
14.5037 | -34.7316 | -24.2439
14.5098 | -34.7321 | -24.2525
14.5159 | -34.7326 | -24.2607
14.522 | -34.7331 | -24.2692
14.5281 | -34.7337 | -24.2775
14.5342 | -34.7342 | -24.286
14.5403 | -34.7347 | -24.2945
14.5465 | -34.7352 | -24.303
14.5526 | -34.7358 | -24.3118
14.5587 | -34.7363 | -24.3203
14.5648 | -34.7369 | -24.3286
14.5709 | -34.7374 | -24.3369
14.577 | -34.7379 | -24.3454
14.5831 | -34.7384 | -24.3538
14.5892 | -34.7389 | -24.3623
14.5953 | -34.7395 | -24.3706
14.6014 | -34.74 | -24.3791
14.6075 | -34.7405 | -24.3874
14.6136 | -34.741 | -24.396
14.6197 | -34.7416 | -24.4047
14.6258 | -34.7421 | -24.4132
14.6319 | -34.7427 | -24.4218
14.638 | -34.7432 | -24.4302
14.6441 | -34.7437 | -24.4387
14.6502 | -34.7442 | -24.4473
14.6563 | -34.7447 | -24.4559
14.6624 | -34.7452 | -24.4644
14.6685 | -34.7458 | -24.4729
14.6746 | -34.7463 | -24.4814
14.6807 | -34.7468 | -24.4898
14.6869 | -34.7474 | -24.4982
14.693 | -34.7479 | -24.5066
14.6991 | -34.7484 | -24.515
14.7052 | -34.749 | -24.5232
14.7113 | -34.7495 | -24.5318
14.7174 | -34.7501 | -24.5402
14.7235 | -34.7506 | -24.5484
14.7296 | -34.7511 | -24.5568
14.7357 | -34.7517 | -24.5653
14.7418 | -34.7522 | -24.5736
14.7479 | -34.7527 | -24.582
14.754 | -34.7532 | -24.5904
14.7601 | -34.7537 | -24.5986
14.7662 | -34.7543 | -24.6069
14.7723 | -34.7548 | -24.6149
14.7784 | -34.7553 | -24.6234
14.7845 | -34.7558 | -24.6319
14.7906 | -34.7564 | -24.6403
14.7967 | -34.7569 | -24.6487
14.8028 | -34.7574 | -24.6572
14.8089 | -34.7579 | -24.6656
14.815 | -34.7585 | -24.6741
14.8211 | -34.759 | -24.6826
14.8272 | -34.7595 | -24.691
14.8334 | -34.7601 | -24.6995
14.8395 | -34.7606 | -24.7077
14.8456 | -34.7611 | -24.7163
14.8517 | -34.7617 | -24.7247
14.8578 | -34.7622 | -24.7331
14.8639 | -34.7628 | -24.7415
14.87 | -34.7633 | -24.75
14.8761 | -34.7638 | -24.7583
14.8822 | -34.7643 | -24.7668
14.8883 | -34.7649 | -24.7753
14.8944 | -34.7653 | -24.7836
14.9005 | -34.7659 | -24.7919
14.9066 | -34.7664 | -24.8006
14.9127 | -34.767 | -24.8088
14.9188 | -34.7675 | -24.8172
14.9249 | -34.768 | -24.8257
14.931 | -34.7685 | -24.8342
14.9371 | -34.7691 | -24.8425
14.9432 | -34.7696 | -24.851
14.9493 | -34.7702 | -24.8592
14.9554 | -34.7707 | -24.8677
14.9615 | -34.7712 | -24.8762
14.9676 | -34.7717 | -24.8848
14.9738 | -34.7723 | -24.8928
14.9799 | -34.7728 | -24.9014
14.986 | -34.7733 | -24.9097
14.9921 | -34.7738 | -24.9179
14.9982 | -34.7744 | -24.9262
15.0043 | -34.7749 | -24.9345
15.0104 | -34.7755 | -24.9431
15.0165 | -34.776 | -24.9517
15.0226 | -34.7766 | -24.9599
15.0287 | -34.7771 | -24.9685
15.0348 | -34.7776 | -24.9769
15.0409 | -34.7782 | -24.9856
15.047 | -34.7787 | -24.9941
15.0531 | -34.7793 | -25.0027
15.0592 | -34.7798 | -25.0112
15.0653 | -34.7804 | -25.0196
15.0714 | -34.7809 | -25.028
15.0775 | -34.7815 | -25.0364
15.0836 | -34.782 | -25.0448
15.0897 | -34.7826 | -25.0531
15.0958 | -34.7831 | -25.0614
15.1019 | -34.7836 | -25.0696
15.108 | -34.7842 | -25.0778
15.1141 | -34.7847 | -25.0862
15.1203 | -34.7852 | -25.0946
15.1264 | -34.7858 | -25.1028
15.1325 | -34.7863 | -25.1112
15.1386 | -34.7868 | -25.1193
15.1447 | -34.7874 | -25.1276
15.1508 | -34.7879 | -25.136
15.1569 | -34.7885 | -25.1443
15.163 | -34.789 | -25.1524
15.1691 | -34.7895 | -25.161
15.1752 | -34.7901 | -25.1693
15.1813 | -34.7906 | -25.1777
15.1874 | -34.7912 | -25.1862
15.1935 | -34.7918 | -25.1943
15.1996 | -34.7923 | -25.2027
15.2057 | -34.7928 | -25.2111
15.2118 | -34.7934 | -25.2195
15.2179 | -34.7939 | -25.2276
15.224 | -34.7945 | -25.2362
15.2301 | -34.795 | -25.2444
15.2362 | -34.7956 | -25.2528
15.2423 | -34.7961 | -25.2613
15.2484 | -34.7967 | -25.2697
15.2545 | -34.7972 | -25.2782
15.2607 | -34.7978 | -25.2866
15.2668 | -34.7984 | -25.2946
15.2729 | -34.7989 | -25.3029
15.279 | -34.7995 | -25.3114
15.2851 | -34.8001 | -25.3198
15.2912 | -34.8006 | -25.3281
15.2973 | -34.8012 | -25.3365
15.3034 | -34.8017 | -25.3448
15.3095 | -34.8023 | -25.3532
15.3156 | -34.8028 | -25.3616
15.3217 | -34.8034 | -25.3698
15.3278 | -34.804 | -25.3783
15.3339 | -34.8045 | -25.3863
15.34 | -34.8051 | -25.3945
15.3461 | -34.8056 | -25.4027
15.3522 | -34.8062 | -25.4109
15.3583 | -34.8067 | -25.4195
15.3644 | -34.8072 | -25.4276
15.3705 | -34.8078 | -25.4357
15.3766 | -34.8083 | -25.4439
15.3827 | -34.8089 | -25.4521
15.3888 | -34.8095 | -25.4602
15.3949 | -34.81 | -25.4686
15.401 | -34.8105 | -25.477
15.4072 | -34.8111 | -25.4851
15.4133 | -34.8116 | -25.4934
15.4194 | -34.8122 | -25.5017
15.4255 | -34.8127 | -25.5098
15.4316 | -34.8133 | -25.5181
15.4377 | -34.8138 | -25.5267
15.4438 | -34.8144 | -25.5348
15.4499 | -34.8149 | -25.5429
15.456 | -34.8155 | -25.5511
15.4621 | -34.816 | -25.5595
15.4682 | -34.8166 | -25.5677
15.4743 | -34.8171 | -25.5758
15.4804 | -34.8177 | -25.584
15.4865 | -34.8182 | -25.5924
15.4926 | -34.8188 | -25.601
15.4987 | -34.8194 | -25.6093
15.5048 | -34.8199 | -25.6179
15.5109 | -34.8204 | -25.6261
15.517 | -34.821 | -25.6344
15.5231 | -34.8215 | -25.6428
15.5292 | -34.8221 | -25.651
15.5353 | -34.8226 | -25.6594
15.5414 | -34.8231 | -25.6678
15.5476 | -34.8237 | -25.6759
15.5537 | -34.8242 | -25.6841
15.5598 | -34.8248 | -25.6926
15.5659 | -34.8253 | -25.7009
15.572 | -34.8259 | -25.7091
15.5781 | -34.8264 | -25.7172
15.5842 | -34.827 | -25.7252
15.5903 | -34.8275 | -25.7334
15.5964 | -34.8281 | -25.7416
15.6025 | -34.8286 | -25.7499
15.6086 | -34.8292 | -25.7581
15.6147 | -34.8297 | -25.7663
15.6208 | -34.8303 | -25.7745
15.6269 | -34.8308 | -25.7827
15.633 | -34.8314 | -25.7908
15.6391 | -34.8319 | -25.799
15.6452 | -34.8325 | -25.8074
15.6513 | -34.833 | -25.8152
15.6574 | -34.8336 | -25.8237
15.6635 | -34.8341 | -25.8318
15.6696 | -34.8347 | -25.8401
15.6757 | -34.8353 | -25.8484
15.6818 | -34.8358 | -25.8566
15.688 | -34.8364 | -25.8646
15.6941 | -34.8369 | -25.8726
15.7002 | -34.8374 | -25.8808
15.7063 | -34.838 | -25.8889
15.7124 | -34.8385 | -25.8973
15.7185 | -34.839 | -25.9054
15.7246 | -34.8396 | -25.9135
15.7307 | -34.8402 | -25.9216
15.7368 | -34.8407 | -25.9299
15.7429 | -34.8413 | -25.938
15.749 | -34.8419 | -25.9464
15.7551 | -34.8424 | -25.9545
15.7612 | -34.843 | -25.9626
15.7673 | -34.8435 | -25.9708
15.7734 | -34.844 | -25.9789
15.7795 | -34.8446 | -25.9873
15.7856 | -34.8451 | -25.9956
15.7917 | -34.8457 | -26.0037
15.7978 | -34.8462 | -26.0121
15.8039 | -34.8468 | -26.0203
15.81 | -34.8473 | -26.0285
15.8161 | -34.8479 | -26.0367
15.8222 | -34.8484 | -26.0446
15.8283 | -34.849 | -26.0525
15.8345 | -34.8495 | -26.0608
15.8406 | -34.8501 | -26.069
15.8467 | -34.8506 | -26.0771
15.8528 | -34.8512 | -26.0855
15.8589 | -34.8518 | -26.0937
15.865 | -34.8524 | -26.1016
15.8711 | -34.8529 | -26.1101
15.8772 | -34.8535 | -26.1183
15.8833 | -34.854 | -26.1266
15.8894 | -34.8545 | -26.1348
15.8955 | -34.855 | -26.143
15.9016 | -34.8556 | -26.151
15.9077 | -34.8561 | -26.1591
15.9138 | -34.8567 | -26.1674
15.9199 | -34.8572 | -26.1755
15.926 | -34.8578 | -26.1836
15.9321 | -34.8583 | -26.1917
15.9382 | -34.8589 | -26.1998
15.9443 | -34.8594 | -26.2078
15.9504 | -34.86 | -26.216
15.9565 | -34.8606 | -26.2241
15.9626 | -34.8611 | -26.2322
15.9687 | -34.8616 | -26.2402
15.9749 | -34.8622 | -26.2483
15.981 | -34.8627 | -26.2564
15.9871 | -34.8632 | -26.2647
15.9932 | -34.8638 | -26.273
15.9993 | -34.8643 | -26.2811
16.0054 | -34.8649 | -26.2889
16.0115 | -34.8654 | -26.297
16.0176 | -34.8659 | -26.3051
16.0237 | -34.8665 | -26.3131
16.0298 | -34.867 | -26.3212
16.0359 | -34.8675 | -26.3294
16.042 | -34.8681 | -26.3373
16.0481 | -34.8687 | -26.3453
16.0542 | -34.8692 | -26.3535
16.0603 | -34.8698 | -26.3615
16.0664 | -34.8703 | -26.3697
16.0725 | -34.8709 | -26.3774
16.0786 | -34.8714 | -26.3853
16.0847 | -34.872 | -26.3935
16.0908 | -34.8725 | -26.4015
16.0969 | -34.873 | -26.4096
16.103 | -34.8736 | -26.4177
16.1091 | -34.8741 | -26.4257
16.1152 | -34.8747 | -26.4338
16.1214 | -34.8753 | -26.4419
16.1275 | -34.8758 | -26.45
16.1336 | -34.8764 | -26.4579
16.1397 | -34.8769 | -26.4662
16.1458 | -34.8775 | -26.474
16.1519 | -34.878 | -26.4822
16.158 | -34.8786 | -26.4903
16.1641 | -34.8791 | -26.4985
16.1702 | -34.8797 | -26.5067
16.1763 | -34.8802 | -26.5146
16.1824 | -34.8808 | -26.5225
16.1885 | -34.8813 | -26.5303
16.1946 | -34.8819 | -26.5382
16.2007 | -34.8824 | -26.5465
16.2068 | -34.883 | -26.5545
16.2129 | -34.8835 | -26.5626
16.219 | -34.8841 | -26.5702
16.2251 | -34.8847 | -26.5782
16.2312 | -34.8852 | -26.5862
16.2373 | -34.8858 | -26.5942
16.2434 | -34.8863 | -26.6026
16.2495 | -34.8869 | -26.6102
16.2556 | -34.8874 | -26.6186
16.2618 | -34.888 | -26.6266
16.2679 | -34.8885 | -26.6347
16.274 | -34.8891 | -26.6428
16.2801 | -34.8896 | -26.6508
16.2862 | -34.8901 | -26.6588
16.2923 | -34.8907 | -26.667
16.2984 | -34.8913 | -26.6751
16.3045 | -34.8917 | -26.6831
16.3106 | -34.8923 | -26.6912
16.3167 | -34.8929 | -26.6991
16.3228 | -34.8934 | -26.7069
16.3289 | -34.8939 | -26.7147
16.335 | -34.8945 | -26.7227
16.3411 | -34.8951 | -26.7308
16.3472 | -34.8956 | -26.7388
16.3533 | -34.8962 | -26.7469
16.3594 | -34.8967 | -26.7547
16.3655 | -34.8973 | -26.7627
16.3716 | -34.8978 | -26.7706
16.3777 | -34.8983 | -26.7785
16.3838 | -34.8989 | -26.7865
16.3899 | -34.8995 | -26.7944
16.396 | -34.9 | -26.8025
16.4021 | -34.9005 | -26.8105
16.4083 | -34.9011 | -26.8181
16.4144 | -34.9016 | -26.8262
16.4205 | -34.9022 | -26.8342
16.4266 | -34.9028 | -26.8418
16.4327 | -34.9033 | -26.8496
16.4388 | -34.9038 | -26.8576
16.4449 | -34.9044 | -26.8657
16.451 | -34.9049 | -26.8737
16.4571 | -34.9055 | -26.8819
16.4632 | -34.906 | -26.8896
16.4693 | -34.9066 | -26.8979
16.4754 | -34.9071 | -26.9063
16.4815 | -34.9076 | -26.9144
16.4876 | -34.9082 | -26.9223
16.4937 | -34.9087 | -26.9302
16.4998 | -34.9093 | -26.9381
16.5059 | -34.9098 | -26.946
16.512 | -34.9104 | -26.9541
16.5181 | -34.9109 | -26.9623
16.5242 | -34.9114 | -26.9703
16.5303 | -34.912 | -26.9783
16.5364 | -34.9125 | -26.9861
16.5425 | -34.913 | -26.9941
16.5487 | -34.9136 | -27.0019
16.5548 | -34.9141 | -27.0099
16.5609 | -34.9147 | -27.0178
16.567 | -34.9152 | -27.0254
16.5731 | -34.9158 | -27.0334
16.5792 | -34.9163 | -27.0411
16.5853 | -34.9168 | -27.049
16.5914 | -34.9174 | -27.057
16.5975 | -34.9179 | -27.0651
16.6036 | -34.9185 | -27.0728
16.6097 | -34.919 | -27.0808
16.6158 | -34.9195 | -27.089
16.6219 | -34.9201 | -27.0968
16.628 | -34.9206 | -27.105
16.6341 | -34.9212 | -27.1129
16.6402 | -34.9217 | -27.1207
16.6463 | -34.9222 | -27.1285
16.6524 | -34.9228 | -27.1365
16.6585 | -34.9233 | -27.1445
16.6646 | -34.9239 | -27.1523
16.6707 | -34.9244 | -27.16
16.6768 | -34.925 | -27.168
16.6829 | -34.9255 | -27.1758
16.689 | -34.9261 | -27.1837
16.6952 | -34.9266 | -27.1916
16.7013 | -34.9271 | -27.1996
16.7074 | -34.9276 | -27.2072
16.7135 | -34.9282 | -27.2151
16.7196 | -34.9287 | -27.223
16.7257 | -34.9293 | -27.2308
16.7318 | -34.9299 | -27.2387
16.7379 | -34.9304 | -27.2466
16.744 | -34.931 | -27.2542
16.7501 | -34.9315 | -27.2622
16.7562 | -34.932 | -27.2701
16.7623 | -34.9326 | -27.2779
16.7684 | -34.9331 | -27.2859
16.7745 | -34.9337 | -27.2938
16.7806 | -34.9342 | -27.3014
16.7867 | -34.9348 | -27.3092
16.7928 | -34.9353 | -27.3171
16.7989 | -34.9359 | -27.325
16.805 | -34.9364 | -27.333
16.8111 | -34.9369 | -27.3408
16.8172 | -34.9375 | -27.3489
16.8233 | -34.9381 | -27.3569
16.8294 | -34.9386 | -27.3647
16.8356 | -34.9391 | -27.3725
16.8417 | -34.9397 | -27.3803
16.8478 | -34.9402 | -27.388
16.8539 | -34.9408 | -27.3958
16.86 | -34.9413 | -27.4037
16.8661 | -34.9418 | -27.4116
16.8722 | -34.9423 | -27.4196
16.8783 | -34.9429 | -27.4277
16.8844 | -34.9434 | -27.4354
16.8905 | -34.944 | -27.4433
16.8966 | -34.9445 | -27.451
16.9027 | -34.945 | -27.4589
16.9088 | -34.9456 | -27.4668
16.9149 | -34.9461 | -27.4747
16.921 | -34.9467 | -27.4823
16.9271 | -34.9472 | -27.4902
16.9332 | -34.9478 | -27.4981
16.9393 | -34.9483 | -27.5061
16.9454 | -34.9488 | -27.5139
16.9515 | -34.9493 | -27.5218
16.9576 | -34.9498 | -27.5296
16.9637 | -34.9504 | -27.5373
16.9698 | -34.951 | -27.5451
16.9759 | -34.9515 | -27.5529
16.9821 | -34.952 | -27.5608
16.9882 | -34.9526 | -27.5682
16.9943 | -34.9531 | -27.576
17.0004 | -34.9536 | -27.5838
17.0065 | -34.9542 | -27.5917
17.0126 | -34.9547 | -27.5996
17.0187 | -34.9553 | -27.6076
17.0248 | -34.9558 | -27.6153
17.0309 | -34.9563 | -27.623
17.037 | -34.9568 | -27.631
17.0431 | -34.9574 | -27.6388
17.0492 | -34.9579 | -27.6466
17.0553 | -34.9584 | -27.6543
17.0614 | -34.959 | -27.6621
17.0675 | -34.9595 | -27.6699
17.0736 | -34.96 | -27.6778
17.0797 | -34.9605 | -27.6855
17.0858 | -34.9611 | -27.6934
17.0919 | -34.9616 | -27.7013
17.098 | -34.9622 | -27.709
17.1041 | -34.9627 | -27.7168
17.1102 | -34.9633 | -27.7246
17.1163 | -34.9639 | -27.7324
17.1225 | -34.9644 | -27.74
17.1286 | -34.965 | -27.7476
17.1347 | -34.9655 | -27.7553
17.1408 | -34.966 | -27.7632
17.1469 | -34.9666 | -27.771
17.153 | -34.9671 | -27.7788
17.1591 | -34.9676 | -27.7866
17.1652 | -34.9682 | -27.7944
17.1713 | -34.9687 | -27.8021
17.1774 | -34.9693 | -27.81
17.1835 | -34.9698 | -27.8176
17.1896 | -34.9704 | -27.8252
17.1957 | -34.9709 | -27.8331
17.2018 | -34.9715 | -27.8407
17.2079 | -34.972 | -27.8485
17.214 | -34.9725 | -27.8564
17.2201 | -34.9731 | -27.8643
17.2262 | -34.9736 | -27.8721
17.2323 | -34.9741 | -27.8799
17.2384 | -34.9747 | -27.8875
17.2445 | -34.9752 | -27.8952
17.2506 | -34.9758 | -27.903
17.2567 | -34.9763 | -27.9107
17.2628 | -34.9769 | -27.9185
17.269 | -34.9774 | -27.9262
17.2751 | -34.9779 | -27.9341
17.2812 | -34.9785 | -27.942
17.2873 | -34.979 | -27.9498
17.2934 | -34.9795 | -27.9576
17.2995 | -34.9801 | -27.9654
17.3056 | -34.9806 | -27.9729
17.3117 | -34.9812 | -27.9807
17.3178 | -34.9817 | -27.9884
17.3239 | -34.9822 | -27.9962
17.33 | -34.9827 | -28.004
17.3361 | -34.9833 | -28.0117
17.3422 | -34.9838 | -28.0191
17.3483 | -34.9843 | -28.0268
17.3544 | -34.9849 | -28.0346
17.3605 | -34.9854 | -28.0423
17.3666 | -34.986 | -28.0501
17.3727 | -34.9865 | -28.0578
17.3788 | -34.987 | -28.0653
17.3849 | -34.9876 | -28.0729
17.391 | -34.9881 | -28.0806
17.3971 | -34.9886 | -28.0883
17.4032 | -34.9891 | -28.0962
17.4094 | -34.9897 | -28.1038
17.4155 | -34.9902 | -28.1114
17.4216 | -34.9907 | -28.1191
17.4277 | -34.9913 | -28.1269
17.4338 | -34.9918 | -28.1347
17.4399 | -34.9923 | -28.1424
17.446 | -34.9929 | -28.1502
17.4521 | -34.9934 | -28.1578
17.4582 | -34.9939 | -28.1653
17.4643 | -34.9944 | -28.1731
17.4704 | -34.995 | -28.1808
17.4765 | -34.9955 | -28.1886
17.4826 | -34.9961 | -28.1961
17.4887 | -34.9966 | -28.2039
17.4948 | -34.9971 | -28.2117
17.5009 | -34.9977 | -28.2194
17.507 | -34.9982 | -28.2272
17.5131 | -34.9987 | -28.235
17.5192 | -34.9993 | -28.2426
17.5253 | -34.9998 | -28.2504
17.5314 | -35.0003 | -28.2581
17.5375 | -35.0009 | -28.2658
17.5436 | -35.0014 | -28.2735
17.5497 | -35.0019 | -28.2812
17.5559 | -35.0025 | -28.2885
17.562 | -35.003 | -28.2961
17.5681 | -35.0036 | -28.3037
17.5742 | -35.0041 | -28.3115
17.5803 | -35.0046 | -28.3192
17.5864 | -35.0052 | -28.3268
17.5925 | -35.0058 | -28.3346
17.5986 | -35.0063 | -28.3422
17.6047 | -35.0068 | -28.35
17.6108 | -35.0074 | -28.3575
17.6169 | -35.0079 | -28.3652
17.623 | -35.0084 | -28.3726
17.6291 | -35.009 | -28.3801
17.6352 | -35.0095 | -28.3878
17.6413 | -35.01 | -28.3954
17.6474 | -35.0106 | -28.4032
17.6535 | -35.0111 | -28.4108
17.6596 | -35.0116 | -28.4183
17.6657 | -35.0122 | -28.426
17.6718 | -35.0127 | -28.4337
17.6779 | -35.0132 | -28.4414
17.684 | -35.0138 | -28.4491
17.6901 | -35.0143 | -28.4568
17.6963 | -35.0148 | -28.4642
17.7024 | -35.0154 | -28.4718
17.7085 | -35.016 | -28.4794
17.7146 | -35.0165 | -28.4871
17.7207 | -35.017 | -28.4947
17.7268 | -35.0176 | -28.5022
17.7329 | -35.0181 | -28.51
17.739 | -35.0186 | -28.5177
17.7451 | -35.0192 | -28.5252
17.7512 | -35.0197 | -28.5327
17.7573 | -35.0203 | -28.5405
17.7634 | -35.0208 | -28.5479
17.7695 | -35.0214 | -28.5555
17.7756 | -35.0219 | -28.5632
17.7817 | -35.0224 | -28.5708
17.7878 | -35.023 | -28.5786
17.7939 | -35.0235 | -28.5863
17.8 | -35.024 | -28.5939
17.8061 | -35.0246 | -28.6015
17.8122 | -35.0251 | -28.6093
17.8183 | -35.0256 | -28.617
17.8244 | -35.0262 | -28.6246
17.8305 | -35.0267 | -28.6323
17.8366 | -35.0272 | -28.6396
17.8428 | -35.0277 | -28.6474
17.8489 | -35.0283 | -28.6549
17.855 | -35.0288 | -28.6625
17.8611 | -35.0293 | -28.67
17.8672 | -35.0299 | -28.6775
17.8733 | -35.0304 | -28.6853
17.8794 | -35.031 | -28.6928
17.8855 | -35.0315 | -28.7004
17.8916 | -35.032 | -28.7082
17.8977 | -35.0326 | -28.7159
17.9038 | -35.0331 | -28.7234
17.9099 | -35.0336 | -28.731
17.916 | -35.0342 | -28.7385
17.9221 | -35.0347 | -28.7462
17.9282 | -35.0352 | -28.7537
17.9343 | -35.0358 | -28.7612
17.9404 | -35.0363 | -28.7686
17.9465 | -35.0368 | -28.7762
17.9526 | -35.0374 | -28.7839
17.9587 | -35.0379 | -28.7915
17.9648 | -35.0385 | -28.7989
17.9709 | -35.039 | -28.8065
17.977 | -35.0395 | -28.814
17.9832 | -35.04 | -28.8217
17.9893 | -35.0406 | -28.8294
17.9954 | -35.0411 | -28.8371
18.0015 | -35.0416 | -28.8447
18.0076 | -35.0422 | -28.8521
18.0137 | -35.0427 | -28.8597
18.0198 | -35.0432 | -28.8672
18.0259 | -35.0437 | -28.8748
18.032 | -35.0442 | -28.8823
18.0381 | -35.0448 | -28.8898
18.0442 | -35.0453 | -28.8973
18.0503 | -35.0458 | -28.9047
18.0564 | -35.0464 | -28.9122
18.0625 | -35.0469 | -28.9199
18.0686 | -35.0474 | -28.9276
18.0747 | -35.048 | -28.9352
18.0808 | -35.0485 | -28.9424
18.0869 | -35.049 | -28.9501
18.093 | -35.0495 | -28.9579
18.0991 | -35.05 | -28.9656
18.1052 | -35.0506 | -28.9733
18.1113 | -35.0511 | -28.981
18.1174 | -35.0517 | -28.9884
18.1236 | -35.0522 | -28.9959
18.1297 | -35.0527 | -29.0035
18.1358 | -35.0533 | -29.0111
18.1419 | -35.0538 | -29.0187
18.148 | -35.0544 | -29.0263
18.1541 | -35.0549 | -29.0334
18.1602 | -35.0555 | -29.0408
18.1663 | -35.056 | -29.0483
18.1724 | -35.0565 | -29.0559
18.1785 | -35.057 | -29.0633
18.1846 | -35.0576 | -29.0707
18.1907 | -35.0581 | -29.0783
18.1968 | -35.0586 | -29.0859
18.2029 | -35.0592 | -29.0936
18.209 | -35.0597 | -29.1012
18.2151 | -35.0603 | -29.1089
18.2212 | -35.0608 | -29.1165
18.2273 | -35.0614 | -29.124
18.2334 | -35.0619 | -29.1315
18.2395 | -35.0624 | -29.1391
18.2456 | -35.0629 | -29.1467
18.2517 | -35.0635 | -29.1541
18.2578 | -35.064 | -29.1615
18.2639 | -35.0646 | -29.1692
18.2701 | -35.0651 | -29.1767
18.2762 | -35.0656 | -29.1842
18.2823 | -35.0661 | -29.1916
18.2884 | -35.0667 | -29.199
18.2945 | -35.0672 | -29.2062
18.3006 | -35.0677 | -29.2136
18.3067 | -35.0683 | -29.2209
18.3128 | -35.0688 | -29.2286
18.3189 | -35.0693 | -29.2362
18.325 | -35.0698 | -29.2432
18.3311 | -35.0703 | -29.2507
18.3372 | -35.0708 | -29.2584
18.3433 | -35.0714 | -29.2658
18.3494 | -35.0719 | -29.2734
18.3555 | -35.0724 | -29.281
18.3616 | -35.073 | -29.2883
18.3677 | -35.0735 | -29.2958
18.3738 | -35.074 | -29.3031
18.3799 | -35.0746 | -29.3105
18.386 | -35.0751 | -29.318
18.3921 | -35.0756 | -29.3256
18.3982 | -35.0762 | -29.3329
18.4043 | -35.0767 | -29.3404
18.4105 | -35.0772 | -29.348
18.4166 | -35.0777 | -29.3553
18.4227 | -35.0782 | -29.3627
18.4288 | -35.0787 | -29.37
18.4349 | -35.0793 | -29.3772
18.441 | -35.0798 | -29.3845
18.4471 | -35.0803 | -29.3917
18.4532 | -35.0808 | -29.3993
18.4593 | -35.0813 | -29.4068
18.4654 | -35.0818 | -29.4141
18.4715 | -35.0824 | -29.4217
18.4776 | -35.0829 | -29.4295
18.4837 | -35.0834 | -29.437
18.4898 | -35.0839 | -29.4446
18.4959 | -35.0845 | -29.4521
18.502 | -35.085 | -29.4593
18.5081 | -35.0855 | -29.4669
18.5142 | -35.086 | -29.4744
18.5203 | -35.0866 | -29.482
18.5264 | -35.0871 | -29.4893
18.5325 | -35.0876 | -29.4967
18.5386 | -35.0882 | -29.504
18.5447 | -35.0887 | -29.5116
18.5508 | -35.0892 | -29.5191
18.557 | -35.0897 | -29.5265
18.5631 | -35.0902 | -29.5339
18.5692 | -35.0907 | -29.5413
18.5753 | -35.0913 | -29.5486
18.5814 | -35.0918 | -29.556
18.5875 | -35.0923 | -29.5631
18.5936 | -35.0928 | -29.5705
18.5997 | -35.0933 | -29.5781
18.6058 | -35.0938 | -29.5853
18.6119 | -35.0944 | -29.5925
18.618 | -35.0949 | -29.6001
18.6241 | -35.0954 | -29.6078
18.6302 | -35.0959 | -29.6153
18.6363 | -35.0964 | -29.6227
18.6424 | -35.0969 | -29.6301
18.6485 | -35.0974 | -29.6377
18.6546 | -35.0979 | -29.6452
18.6607 | -35.0985 | -29.6529
18.6668 | -35.099 | -29.6601
18.6729 | -35.0995 | -29.6676
18.679 | -35.1 | -29.6748
18.6851 | -35.1006 | -29.6821
18.6912 | -35.1011 | -29.6895
18.6974 | -35.1016 | -29.6967
18.7035 | -35.1021 | -29.7041
18.7096 | -35.1026 | -29.7116
18.7157 | -35.1031 | -29.719
18.7218 | -35.1036 | -29.7263
18.7279 | -35.1042 | -29.7337
18.734 | -35.1047 | -29.7412
18.7401 | -35.1053 | -29.7487
18.7462 | -35.1058 | -29.7559
18.7523 | -35.1063 | -29.7633
18.7584 | -35.1068 | -29.7709
18.7645 | -35.1074 | -29.7784
18.7706 | -35.1079 | -29.7857
18.7767 | -35.1084 | -29.7932
18.7828 | -35.1089 | -29.8006
18.7889 | -35.1095 | -29.808
18.795 | -35.1099 | -29.8156
18.8011 | -35.1104 | -29.823
18.8072 | -35.111 | -29.8304
18.8133 | -35.1115 | -29.8379
18.8194 | -35.112 | -29.8452
18.8255 | -35.1125 | -29.8526
18.8316 | -35.113 | -29.86
18.8377 | -35.1135 | -29.8673
18.8439 | -35.114 | -29.8746
18.85 | -35.1145 | -29.8823
18.8561 | -35.115 | -29.8896
18.8622 | -35.1156 | -29.8969
18.8683 | -35.1161 | -29.9042
18.8744 | -35.1166 | -29.9114
18.8805 | -35.1172 | -29.9187
18.8866 | -35.1177 | -29.9261
18.8927 | -35.1182 | -29.9332
18.8988 | -35.1187 | -29.9405
18.9049 | -35.1193 | -29.9479
18.911 | -35.1198 | -29.9553
18.9171 | -35.1203 | -29.9628
18.9232 | -35.1208 | -29.9699
18.9293 | -35.1213 | -29.9773
18.9354 | -35.1219 | -29.9847
18.9415 | -35.1224 | -29.9918
18.9476 | -35.1229 | -29.9992
18.9537 | -35.1234 | -30.0066
18.9598 | -35.1239 | -30.0138
18.9659 | -35.1244 | -30.0211
18.972 | -35.1249 | -30.0284
18.9781 | -35.1254 | -30.0359
18.9843 | -35.1259 | -30.0433
18.9904 | -35.1264 | -30.0506
18.9965 | -35.127 | -30.0576
19.0026 | -35.1275 | -30.0649
19.0087 | -35.128 | -30.0721
19.0148 | -35.1285 | -30.0794
19.0209 | -35.129 | -30.087
19.027 | -35.1295 | -30.0943
19.0331 | -35.1301 | -30.1016
19.0392 | -35.1306 | -30.1091
19.0453 | -35.1311 | -30.1165
19.0514 | -35.1316 | -30.1238
19.0575 | -35.1322 | -30.1312
19.0636 | -35.1327 | -30.1383
19.0697 | -35.1332 | -30.1457
19.0758 | -35.1337 | -30.1529
19.0819 | -35.1342 | -30.1601
19.088 | -35.1348 | -30.1674
19.0941 | -35.1353 | -30.175
19.1002 | -35.1358 | -30.1822
19.1063 | -35.1363 | -30.1896
19.1124 | -35.1368 | -30.1969
19.1185 | -35.1373 | -30.2041
19.1246 | -35.1378 | -30.2113
19.1308 | -35.1383 | -30.2188
19.1369 | -35.1388 | -30.2258
19.143 | -35.1394 | -30.2335
19.1491 | -35.1399 | -30.2409
19.1552 | -35.1404 | -30.2482
19.1613 | -35.1409 | -30.2555
19.1674 | -35.1414 | -30.2629
19.1735 | -35.1419 | -30.2702
19.1796 | -35.1424 | -30.2774
19.1857 | -35.1429 | -30.2847
19.1918 | -35.1434 | -30.292
19.1979 | -35.1439 | -30.2994
19.204 | -35.1445 | -30.3067
19.2101 | -35.145 | -30.314
19.2162 | -35.1455 | -30.3212
19.2223 | -35.146 | -30.3286
19.2284 | -35.1465 | -30.336
19.2345 | -35.1469 | -30.3432
19.2406 | -35.1474 | -30.3504
19.2467 | -35.148 | -30.3577
19.2528 | -35.1485 | -30.365
19.2589 | -35.149 | -30.3721
19.265 | -35.1495 | -30.3795
19.2712 | -35.15 | -30.3867
19.2773 | -35.1505 | -30.3941
19.2834 | -35.1511 | -30.4013
19.2895 | -35.1516 | -30.4086
19.2956 | -35.1521 | -30.4159
19.3017 | -35.1526 | -30.4233
19.3078 | -35.1531 | -30.4306
19.3139 | -35.1536 | -30.4376
19.32 | -35.1541 | -30.4448
19.3261 | -35.1546 | -30.452
19.3322 | -35.1551 | -30.4592
19.3383 | -35.1557 | -30.4664
19.3444 | -35.1562 | -30.4735
19.3505 | -35.1567 | -30.4808
19.3566 | -35.1572 | -30.4882
19.3627 | -35.1577 | -30.4955
19.3688 | -35.1583 | -30.5026
19.3749 | -35.1588 | -30.5102
19.381 | -35.1593 | -30.5172
19.3871 | -35.1598 | -30.5245
19.3932 | -35.1603 | -30.5318
19.3993 | -35.1609 | -30.5391
19.4054 | -35.1614 | -30.5465
19.4115 | -35.1619 | -30.554
19.4177 | -35.1624 | -30.5613
19.4238 | -35.1629 | -30.5684
19.4299 | -35.1635 | -30.5759
19.436 | -35.164 | -30.5835
19.4421 | -35.1645 | -30.5905
19.4482 | -35.165 | -30.5977
19.4543 | -35.1655 | -30.6049
19.4604 | -35.1661 | -30.6122
19.4665 | -35.1666 | -30.6192
19.4726 | -35.1671 | -30.6264
19.4787 | -35.1676 | -30.6334
19.4848 | -35.1681 | -30.6408
19.4909 | -35.1686 | -30.648
19.497 | -35.1691 | -30.6552
19.5031 | -35.1696 | -30.6624
19.5092 | -35.1702 | -30.6696
19.5153 | -35.1707 | -30.6768
19.5214 | -35.1712 | -30.684
19.5275 | -35.1717 | -30.6913
19.5336 | -35.1721 | -30.6986
19.5397 | -35.1726 | -30.7059
19.5458 | -35.1732 | -30.7135
19.5519 | -35.1737 | -30.7207
19.5581 | -35.1741 | -30.7277
19.5642 | -35.1746 | -30.7351
19.5703 | -35.175 | -30.7425
19.5764 | -35.1756 | -30.7498
19.5825 | -35.176 | -30.757
19.5886 | -35.1766 | -30.7643
19.5947 | -35.1771 | -30.7714
19.6008 | -35.1776 | -30.7782
19.6069 | -35.1781 | -30.7856
19.613 | -35.1786 | -30.7928
19.6191 | -35.1791 | -30.8002
19.6252 | -35.1796 | -30.8075
19.6313 | -35.1802 | -30.8148
19.6374 | -35.1806 | -30.8223
19.6435 | -35.1812 | -30.8294
19.6496 | -35.1817 | -30.8363
19.6557 | -35.1822 | -30.8434
19.6618 | -35.1827 | -30.8505
19.6679 | -35.1832 | -30.8579
19.674 | -35.1837 | -30.8653
19.6801 | -35.1843 | -30.8726
19.6862 | -35.1847 | -30.8799
19.6923 | -35.1853 | -30.8871
19.6984 | -35.1857 | -30.8942
19.7046 | -35.1862 | -30.9014
19.7107 | -35.1867 | -30.9086
19.7168 | -35.1872 | -30.916
19.7229 | -35.1877 | -30.9232
19.729 | -35.1882 | -30.9304
19.7351 | -35.1887 | -30.9374
19.7412 | -35.1892 | -30.9445
19.7473 | -35.1897 | -30.9515
19.7534 | -35.1903 | -30.9588
19.7595 | -35.1908 | -30.9662
19.7656 | -35.1913 | -30.9735
19.7717 | -35.1918 | -30.9799
19.7778 | -35.1923 | -30.987
19.7839 | -35.1928 | -30.9942
19.79 | -35.1934 | -31.0014
19.7961 | -35.1938 | -31.0085
19.8022 | -35.1944 | -31.0155
19.8083 | -35.195 | -31.0227
19.8144 | -35.1955 | -31.0299
19.8205 | -35.196 | -31.0371
19.8266 | -35.1965 | -31.0442
19.8327 | -35.197 | -31.0512
19.8388 | -35.1975 | -31.058
19.845 | -35.198 | -31.0651
19.8511 | -35.1985 | -31.0726
19.8572 | -35.199 | -31.0797
19.8633 | -35.1995 | -31.087
19.8694 | -35.2 | -31.0945
19.8755 | -35.2005 | -31.1013
19.8816 | -35.201 | -31.1085
19.8877 | -35.2015 | -31.1161
19.8938 | -35.2019 | -31.1231
19.8999 | -35.2024 | -31.1303
19.906 | -35.2029 | -31.1376
19.9121 | -35.2034 | -31.1447
19.9182 | -35.2039 | -31.152
19.9243 | -35.2045 | -31.1591
19.9304 | -35.2049 | -31.1662
19.9365 | -35.2054 | -31.1735
19.9426 | -35.2059 | -31.1805
19.9487 | -35.2064 | -31.1877
19.9548 | -35.2069 | -31.195
19.9609 | -35.2075 | -31.2021
19.967 | -35.208 | -31.2091
19.9731 | -35.2085 | -31.2161
19.9792 | -35.209 | -31.2231
19.9853 | -35.2095 | -31.23
19.9915 | -35.21 | -31.2369
19.9976 | -35.2105 | -31.2438
20.0037 | -35.211 | -31.2507
20.0098 | -35.2115 | -31.2578
20.0159 | -35.212 | -31.2647
20.022 | -35.2125 | -31.2721
20.0281 | -35.213 | -31.2792
20.0342 | -35.2135 | -31.2863
20.0403 | -35.214 | -31.2935
20.0464 | -35.2145 | -31.301
20.0525 | -35.215 | -31.3076
20.0586 | -35.2156 | -31.315
20.0647 | -35.2161 | -31.3223
20.0708 | -35.2166 | -31.3298
20.0769 | -35.2171 | -31.337
20.083 | -35.2176 | -31.3437
20.0891 | -35.2181 | -31.3509
20.0952 | -35.2186 | -31.3581
20.1013 | -35.2191 | -31.3651
20.1074 | -35.2196 | -31.3724
20.1135 | -35.2201 | -31.3794
20.1196 | -35.2205 | -31.3864
20.1257 | -35.221 | -31.3934
20.1319 | -35.2216 | -31.4003
20.138 | -35.2221 | -31.4074
20.1441 | -35.2226 | -31.4144
20.1502 | -35.2231 | -31.4215
20.1563 | -35.2236 | -31.4285
20.1624 | -35.2241 | -31.4355
20.1685 | -35.2247 | -31.4427
20.1746 | -35.2252 | -31.4495
20.1807 | -35.2257 | -31.4566
20.1868 | -35.2262 | -31.4639
20.1929 | -35.2267 | -31.4707
20.199 | -35.2272 | -31.478
20.2051 | -35.2278 | -31.4852
20.2112 | -35.2283 | -31.4922
20.2173 | -35.2289 | -31.4992
20.2234 | -35.2294 | -31.5065
20.2295 | -35.23 | -31.5141
20.2356 | -35.2305 | -31.521
20.2417 | -35.231 | -31.5282
20.2478 | -35.2315 | -31.5351
20.2539 | -35.232 | -31.5422
20.26 | -35.2325 | -31.5495
20.2661 | -35.2331 | -31.5569
20.2723 | -35.2336 | -31.564
20.2784 | -35.2341 | -31.5713
20.2845 | -35.2346 | -31.5786
20.2906 | -35.2351 | -31.5858
20.2967 | -35.2356 | -31.593
20.3028 | -35.2361 | -31.5999
20.3089 | -35.2366 | -31.6069
20.315 | -35.2371 | -31.6143
20.3211 | -35.2377 | -31.6215
20.3272 | -35.2383 | -31.6287
20.3333 | -35.2388 | -31.6357
20.3394 | -35.2393 | -31.6427
20.3455 | -35.2398 | -31.6498
20.3516 | -35.2403 | -31.657
20.3577 | -35.2408 | -31.6639
20.3638 | -35.2412 | -31.6707
20.3699 | -35.2418 | -31.6777
20.376 | -35.2423 | -31.6849
20.3821 | -35.2428 | -31.6922
20.3882 | -35.2434 | -31.6992
20.3943 | -35.2439 | -31.7063
20.4004 | -35.2444 | -31.7132
20.4065 | -35.2449 | -31.7201
20.4126 | -35.2454 | -31.7269
20.4188 | -35.2459 | -31.7342
20.4249 | -35.2464 | -31.7409
20.431 | -35.2469 | -31.7479
20.4371 | -35.2474 | -31.7549
20.4432 | -35.2479 | -31.7623
20.4493 | -35.2485 | -31.7697
20.4554 | -35.249 | -31.7767
20.4615 | -35.2495 | -31.7838
20.4676 | -35.25 | -31.7906
20.4737 | -35.2504 | -31.7981
20.4798 | -35.251 | -31.8051
20.4859 | -35.2514 | -31.8124
20.492 | -35.252 | -31.8193
20.4981 | -35.2525 | -31.8263
20.5042 | -35.253 | -31.8334
20.5103 | -35.2535 | -31.8402
20.5164 | -35.254 | -31.847
20.5225 | -35.2545 | -31.8544
20.5286 | -35.255 | -31.8616
20.5347 | -35.2556 | -31.8688
20.5408 | -35.2561 | -31.8752
20.5469 | -35.2566 | -31.8824
20.553 | -35.2572 | -31.8897
20.5592 | -35.2576 | -31.8968
20.5653 | -35.2581 | -31.9039
20.5714 | -35.2587 | -31.9114
20.5775 | -35.2592 | -31.9185
20.5836 | -35.2597 | -31.9255
20.5897 | -35.2602 | -31.9324
20.5958 | -35.2606 | -31.9397
20.6019 | -35.2611 | -31.9469
20.608 | -35.2616 | -31.954
20.6141 | -35.2622 | -31.961
20.6202 | -35.2627 | -31.968
20.6263 | -35.2632 | -31.9752
20.6324 | -35.2638 | -31.9823
20.6385 | -35.2643 | -31.9893
20.6446 | -35.2647 | -31.9964
20.6507 | -35.2652 | -32.0032
20.6568 | -35.2657 | -32.0103
20.6629 | -35.2662 | -32.0174
20.669 | -35.2667 | -32.0247
20.6751 | -35.2672 | -32.0317
20.6812 | -35.2678 | -32.0383
20.6873 | -35.2683 | -32.0454
20.6934 | -35.2688 | -32.0525
20.6995 | -35.2693 | -32.0593
20.7057 | -35.2697 | -32.0664
20.7118 | -35.2703 | -32.0738
20.7179 | -35.2708 | -32.0804
20.724 | -35.2714 | -32.0876
20.7301 | -35.2719 | -32.0945
20.7362 | -35.2724 | -32.1015
20.7423 | -35.2729 | -32.1084
20.7484 | -35.2735 | -32.1154
20.7545 | -35.274 | -32.1222
20.7606 | -35.2746 | -32.1295
20.7667 | -35.2751 | -32.1364
20.7728 | -35.2756 | -32.1434
20.7789 | -35.2761 | -32.1504
20.785 | -35.2766 | -32.1573
20.7911 | -35.2772 | -32.1646
20.7972 | -35.2776 | -32.1713
20.8033 | -35.2781 | -32.1785
20.8094 | -35.2787 | -32.185
20.8155 | -35.2793 | -32.192
20.8216 | -35.2797 | -32.199
20.8277 | -35.2803 | -32.2057
20.8338 | -35.2809 | -32.2125
20.8399 | -35.2815 | -32.2196
20.8461 | -35.282 | -32.2266
20.8522 | -35.2825 | -32.2334
20.8583 | -35.2831 | -32.2404
20.8644 | -35.2836 | -32.2476
20.8705 | -35.2842 | -32.2548
20.8766 | -35.2847 | -32.2623
20.8827 | -35.2852 | -32.2693
20.8888 | -35.2858 | -32.2762
20.8949 | -35.2863 | -32.2835
20.901 | -35.2869 | -32.2907
20.9071 | -35.2875 | -32.2972
20.9132 | -35.288 | -32.3041
20.9193 | -35.2886 | -32.3113
20.9254 | -35.2891 | -32.3182
20.9315 | -35.2896 | -32.3251
20.9376 | -35.2901 | -32.3321
20.9437 | -35.2906 | -32.3391
20.9498 | -35.2911 | -32.3466
20.9559 | -35.2916 | -32.3538
20.962 | -35.2922 | -32.3608
20.9681 | -35.2928 | -32.3674
20.9742 | -35.2933 | -32.3747
20.9803 | -35.2938 | -32.382
20.9864 | -35.2943 | -32.3893
20.9926 | -35.2948 | -32.3961
20.9987 | -35.2954 | -32.4029
21.0048 | -35.296 | -32.4095
21.0109 | -35.2965 | -32.4165
21.017 | -35.2971 | -32.4235
21.0231 | -35.2977 | -32.4311
21.0292 | -35.2982 | -32.4383
21.0353 | -35.2987 | -32.445
21.0414 | -35.2992 | -32.4517
21.0475 | -35.2998 | -32.4584
21.0536 | -35.3003 | -32.4654
21.0597 | -35.3008 | -32.4725
21.0658 | -35.3014 | -32.4799
21.0719 | -35.3021 | -32.4863
21.078 | -35.3026 | -32.4933
21.0841 | -35.3032 | -32.5007
21.0902 | -35.3038 | -32.5072
21.0963 | -35.3042 | -32.5139
21.1024 | -35.3047 | -32.5212
21.1085 | -35.3052 | -32.5271
21.1146 | -35.3057 | -32.5338
21.1207 | -35.3063 | -32.5411
21.1268 | -35.3068 | -32.548
21.133 | -35.3073 | -32.5549
21.1391 | -35.3078 | -32.5617
21.1452 | -35.3083 | -32.5684
21.1513 | -35.3089 | -32.5756
21.1574 | -35.3094 | -32.5825
21.1635 | -35.3099 | -32.5895
21.1696 | -35.3103 | -32.5967
21.1757 | -35.3108 | -32.6037
21.1818 | -35.3114 | -32.6102
21.1879 | -35.3119 | -32.6171
21.194 | -35.3124 | -32.6243
21.2001 | -35.3129 | -32.6309
21.2062 | -35.3135 | -32.6382
21.2123 | -35.3141 | -32.6442
21.2184 | -35.3145 | -32.6517
21.2245 | -35.315 | -32.6588
21.2306 | -35.3156 | -32.6656
21.2367 | -35.3161 | -32.6722
21.2428 | -35.3167 | -32.6791
21.2489 | -35.3173 | -32.6853
21.255 | -35.3178 | -32.692
21.2611 | -35.3184 | -32.6987
21.2672 | -35.319 | -32.7052
21.2733 | -35.3195 | -32.7126
21.2795 | -35.3202 | -32.7196
21.2856 | -35.3207 | -32.7261
21.2917 | -35.3213 | -32.7333
21.2978 | -35.3218 | -32.7405
21.3039 | -35.3223 | -32.7477
21.31 | -35.3228 | -32.7547
21.3161 | -35.3234 | -32.761
21.3222 | -35.3239 | -32.7682
21.3283 | -35.3244 | -32.7753
21.3344 | -35.3249 | -32.7827
21.3405 | -35.3255 | -32.7895
21.3466 | -35.326 | -32.7963
21.3527 | -35.3265 | -32.8027
21.3588 | -35.3271 | -32.8095
21.3649 | -35.3277 | -32.8161
21.371 | -35.3282 | -32.8232
21.3771 | -35.3288 | -32.83
21.3832 | -35.3292 | -32.8367
21.3893 | -35.3299 | -32.843
21.3954 | -35.3304 | -32.8504
21.4015 | -35.3309 | -32.8569
21.4076 | -35.3315 | -32.8639
21.4137 | -35.332 | -32.8704
21.4199 | -35.3327 | -32.8773
21.426 | -35.3331 | -32.8835
21.4321 | -35.3337 | -32.8903
21.4382 | -35.3343 | -32.8973
21.4443 | -35.3348 | -32.9043
21.4504 | -35.3353 | -32.9113
21.4565 | -35.3359 | -32.9182
21.4626 | -35.3365 | -32.9255
21.4687 | -35.337 | -32.9322
21.4748 | -35.3376 | -32.939
21.4809 | -35.3382 | -32.9461
21.487 | -35.3387 | -32.9532
21.4931 | -35.3393 | -32.9599
21.4992 | -35.3398 | -32.9668
21.5053 | -35.3404 | -32.9738
21.5114 | -35.341 | -32.981
21.5175 | -35.3415 | -32.988
21.5236 | -35.342 | -32.9949
21.5297 | -35.3425 | -33.0015
21.5358 | -35.3431 | -33.0087
21.5419 | -35.3436 | -33.0158
21.548 | -35.3441 | -33.0225
21.5541 | -35.3447 | -33.0301
21.5602 | -35.3452 | -33.037
21.5664 | -35.3458 | -33.0438
21.5725 | -35.3463 | -33.0505
21.5786 | -35.3468 | -33.0575
21.5847 | -35.3473 | -33.0646
21.5908 | -35.3479 | -33.0711
21.5969 | -35.3485 | -33.0784
21.603 | -35.3491 | -33.0853
21.6091 | -35.3497 | -33.0923
21.6152 | -35.3503 | -33.0994
21.6213 | -35.3508 | -33.1065
21.6274 | -35.3514 | -33.1134
21.6335 | -35.3519 | -33.1206
21.6396 | -35.3525 | -33.1279
21.6457 | -35.353 | -33.1353
21.6518 | -35.3536 | -33.142
21.6579 | -35.3541 | -33.1491
21.664 | -35.3546 | -33.156
21.6701 | -35.3552 | -33.1627
21.6762 | -35.3558 | -33.1694
21.6823 | -35.3564 | -33.176
21.6884 | -35.3568 | -33.183
21.6945 | -35.3574 | -33.1898
21.7006 | -35.3579 | -33.1967
21.7068 | -35.3585 | -33.2033
21.7129 | -35.359 | -33.2102
21.719 | -35.3595 | -33.2171
21.7251 | -35.3601 | -33.2243
21.7312 | -35.3606 | -33.2312
21.7373 | -35.3612 | -33.2379
21.7434 | -35.3617 | -33.2447
21.7495 | -35.3623 | -33.2513
21.7556 | -35.3629 | -33.2582
21.7617 | -35.3634 | -33.2653
21.7678 | -35.364 | -33.2724
21.7739 | -35.3647 | -33.2788
21.78 | -35.3653 | -33.2853
21.7861 | -35.366 | -33.2921
21.7922 | -35.3665 | -33.2993
21.7983 | -35.367 | -33.3063
21.8044 | -35.3675 | -33.313
21.8105 | -35.3682 | -33.3196
21.8166 | -35.3688 | -33.3263
21.8227 | -35.3694 | -33.3328
21.8288 | -35.3699 | -33.3395
21.8349 | -35.3705 | -33.3463
21.841 | -35.371 | -33.3531
21.8471 | -35.3716 | -33.3593
21.8533 | -35.3722 | -33.3667
21.8594 | -35.3729 | -33.373
21.8655 | -35.3734 | -33.38
21.8716 | -35.3738 | -33.3868
21.8777 | -35.3744 | -33.393
21.8838 | -35.3751 | -33.4001
21.8899 | -35.3757 | -33.4071
21.896 | -35.3762 | -33.4134
21.9021 | -35.3769 | -33.4199
21.9082 | -35.3774 | -33.4277
21.9143 | -35.3779 | -33.4346
21.9204 | -35.3785 | -33.4417
21.9265 | -35.379 | -33.4484
21.9326 | -35.3796 | -33.4553
21.9387 | -35.3801 | -33.4619
21.9448 | -35.3806 | -33.4693
21.9509 | -35.3811 | -33.4752
21.957 | -35.3816 | -33.4814
21.9631 | -35.3821 | -33.4886
21.9692 | -35.3826 | -33.4953
21.9753 | -35.3832 | -33.5025
21.9814 | -35.3837 | -33.5094
21.9875 | -35.3844 | -33.5156
21.9937 | -35.3851 | -33.5225
21.9998 | -35.3856 | -33.5303
22.0059 | -35.3861 | -33.5373
22.012 | -35.3866 | -33.5446
22.0181 | -35.3873 | -33.5517
22.0242 | -35.3879 | -33.5586
22.0303 | -35.3885 | -33.5654
22.0364 | -35.389 | -33.5721
22.0425 | -35.3897 | -33.5789
22.0486 | -35.3904 | -33.5857
22.0547 | -35.3909 | -33.5916
22.0608 | -35.3914 | -33.5984
22.0669 | -35.392 | -33.6054
22.073 | -35.3925 | -33.612
22.0791 | -35.3931 | -33.6189
22.0852 | -35.3937 | -33.6255
22.0913 | -35.3944 | -33.6321
22.0974 | -35.3948 | -33.6388
22.1035 | -35.3953 | -33.6451
22.1096 | -35.3958 | -33.6513
22.1157 | -35.3964 | -33.6587
22.1218 | -35.397 | -33.6652
22.1279 | -35.3976 | -33.6721
22.134 | -35.3982 | -33.6791
22.1402 | -35.3988 | -33.6862
22.1463 | -35.3994 | -33.6934
22.1524 | -35.3999 | -33.7
22.1585 | -35.4006 | -33.7063
22.1646 | -35.4012 | -33.7129
22.1707 | -35.4017 | -33.7197
22.1768 | -35.4022 | -33.7266
22.1829 | -35.4028 | -33.733
22.189 | -35.4034 | -33.7398
22.1951 | -35.4039 | -33.7467
22.2012 | -35.4043 | -33.7531
22.2073 | -35.4049 | -33.7598
22.2134 | -35.4055 | -33.7662
22.2195 | -35.4061 | -33.7731
22.2256 | -35.4067 | -33.7799
22.2317 | -35.4074 | -33.7867
22.2378 | -35.4078 | -33.7933
22.2439 | -35.4084 | -33.8001
22.25 | -35.409 | -33.807
22.2561 | -35.4096 | -33.814
22.2622 | -35.4102 | -33.8205
22.2683 | -35.4108 | -33.8268
22.2744 | -35.4114 | -33.8335
22.2806 | -35.4119 | -33.8401
22.2867 | -35.4125 | -33.847
22.2928 | -35.4131 | -33.8537
22.2989 | -35.4136 | -33.86
22.305 | -35.4142 | -33.8668
22.3111 | -35.4148 | -33.8738
22.3172 | -35.4154 | -33.8809
22.3233 | -35.416 | -33.8876
22.3294 | -35.4166 | -33.8942
22.3355 | -35.4171 | -33.901
22.3416 | -35.4177 | -33.9078
22.3477 | -35.4182 | -33.9139
22.3538 | -35.4188 | -33.9211
22.3599 | -35.4193 | -33.9278
22.366 | -35.4199 | -33.9337
22.3721 | -35.4204 | -33.9399
22.3782 | -35.4209 | -33.9472
22.3843 | -35.4216 | -33.9536
22.3904 | -35.422 | -33.9601
22.3965 | -35.4227 | -33.9675
22.4026 | -35.4233 | -33.9744
22.4087 | -35.4242 | -33.9819
22.4148 | -35.4246 | -33.9886
22.4209 | -35.4253 | -33.9955
22.4271 | -35.4258 | -34.0016
22.4332 | -35.4264 | -34.0081
22.4393 | -35.4269 | -34.0152
22.4454 | -35.4276 | -34.0218
22.4515 | -35.4282 | -34.0284
22.4576 | -35.4288 | -34.0351
22.4637 | -35.4294 | -34.0423
22.4698 | -35.43 | -34.0484
22.4759 | -35.4306 | -34.0539
22.482 | -35.4312 | -34.0604
22.4881 | -35.4318 | -34.0672
22.4942 | -35.4324 | -34.0741
22.5003 | -35.433 | -34.0811
22.5064 | -35.4336 | -34.0873
22.5125 | -35.4342 | -34.0929
22.5186 | -35.4347 | -34.0999
22.5247 | -35.4353 | -34.1071
22.5308 | -35.436 | -34.1142
22.5369 | -35.4365 | -34.1211
22.543 | -35.4371 | -34.1277
22.5491 | -35.4377 | -34.1342
22.5552 | -35.4382 | -34.1407
22.5613 | -35.4389 | -34.1468
22.5675 | -35.4395 | -34.1533
22.5736 | -35.4401 | -34.1603
22.5797 | -35.4407 | -34.1668
22.5858 | -35.4413 | -34.173
22.5919 | -35.4419 | -34.1798
22.598 | -35.4425 | -34.186
22.6041 | -35.443 | -34.1928
22.6102 | -35.4436 | -34.1999
22.6163 | -35.4443 | -34.2066
22.6224 | -35.4449 | -34.2129
22.6285 | -35.4455 | -34.2193
22.6346 | -35.4461 | -34.2257
22.6407 | -35.4466 | -34.2324
22.6468 | -35.4472 | -34.2389
22.6529 | -35.4479 | -34.2455
22.659 | -35.4484 | -34.2518
22.6651 | -35.4489 | -34.2586
22.6712 | -35.4496 | -34.2651
22.6773 | -35.4502 | -34.2716
22.6834 | -35.4508 | -34.2783
22.6895 | -35.4514 | -34.2848
22.6956 | -35.4519 | -34.2911
22.7017 | -35.4526 | -34.2977
22.7079 | -35.4532 | -34.3046
22.714 | -35.4537 | -34.3108
22.7201 | -35.4543 | -34.3172
22.7262 | -35.4549 | -34.3239
22.7323 | -35.4555 | -34.3306
22.7384 | -35.4561 | -34.3375
22.7445 | -35.4566 | -34.3437
22.7506 | -35.4572 | -34.3508
22.7567 | -35.4578 | -34.3571
22.7628 | -35.4585 | -34.3632
22.7689 | -35.459 | -34.3697
22.775 | -35.4596 | -34.3769
22.7811 | -35.4602 | -34.3832
22.7872 | -35.4608 | -34.3898
22.7933 | -35.4614 | -34.396
22.7994 | -35.4619 | -34.4027
22.8055 | -35.4625 | -34.4091
22.8116 | -35.4631 | -34.4158
22.8177 | -35.4636 | -34.4222
22.8238 | -35.4642 | -34.4288
22.8299 | -35.4648 | -34.435
22.836 | -35.4654 | -34.4413
22.8421 | -35.466 | -34.4478
22.8482 | -35.4665 | -34.4548
22.8544 | -35.4671 | -34.4618
22.8605 | -35.4676 | -34.4683
22.8666 | -35.4682 | -34.4746
22.8727 | -35.4688 | -34.4814
22.8788 | -35.4694 | -34.4882
22.8849 | -35.4699 | -34.4948
22.891 | -35.4705 | -34.5011
22.8971 | -35.4711 | -34.5075
22.9032 | -35.4717 | -34.5135
22.9093 | -35.4723 | -34.5199
22.9154 | -35.4728 | -34.5266
22.9215 | -35.4734 | -34.5336
22.9276 | -35.474 | -34.5402
22.9337 | -35.4746 | -34.5463
22.9398 | -35.4752 | -34.5529
22.9459 | -35.4758 | -34.5598
22.952 | -35.4764 | -34.5667
22.9581 | -35.4769 | -34.5732
22.9642 | -35.4775 | -34.5797
22.9703 | -35.4782 | -34.5855
22.9764 | -35.4788 | -34.5923
22.9825 | -35.4794 | -34.5987
22.9886 | -35.48 | -34.6054
22.9948 | -35.4805 | -34.6117
23.0009 | -35.4811 | -34.6187
23.007 | -35.4818 | -34.6255
23.0131 | -35.4824 | -34.6323
23.0192 | -35.4829 | -34.6384
23.0253 | -35.4835 | -34.6452
23.0314 | -35.4841 | -34.6517
23.0375 | -35.4847 | -34.6578
23.0436 | -35.4854 | -34.6643
23.0497 | -35.4859 | -34.6708
23.0558 | -35.4865 | -34.6776
23.0619 | -35.4871 | -34.6838
23.068 | -35.4876 | -34.69
23.0741 | -35.4882 | -34.6958
23.0802 | -35.4889 | -34.7017
23.0863 | -35.4895 | -34.7082
23.0924 | -35.49 | -34.7148
23.0985 | -35.4906 | -34.7213
23.1046 | -35.4912 | -34.7277
23.1107 | -35.4918 | -34.7342
23.1168 | -35.4923 | -34.7405
23.1229 | -35.4929 | -34.7471
23.129 | -35.4935 | -34.7539
23.1351 | -35.494 | -34.7604
23.1413 | -35.4945 | -34.7666
23.1474 | -35.4951 | -34.7728
23.1535 | -35.4956 | -34.7791
23.1596 | -35.4962 | -34.7853
23.1657 | -35.4968 | -34.7915
23.1718 | -35.4974 | -34.7978
23.1779 | -35.498 | -34.8046
23.184 | -35.4985 | -34.8108
23.1901 | -35.4991 | -34.8172
23.1962 | -35.4998 | -34.8238
23.2023 | -35.5004 | -34.8303
23.2084 | -35.501 | -34.837
23.2145 | -35.5016 | -34.8434
23.2206 | -35.5021 | -34.8497
23.2267 | -35.5027 | -34.8565
23.2328 | -35.5033 | -34.8625
23.2389 | -35.5039 | -34.8689
23.245 | -35.5045 | -34.8751
23.2511 | -35.5051 | -34.8815
23.2572 | -35.5056 | -34.8881
23.2633 | -35.5062 | -34.8941
23.2694 | -35.5067 | -34.9006
23.2755 | -35.5073 | -34.9067
23.2817 | -35.5079 | -34.9131
23.2878 | -35.5085 | -34.9189
23.2939 | -35.509 | -34.9255
23.3 | -35.5096 | -34.9317
23.3061 | -35.5102 | -34.9381
23.3122 | -35.5108 | -34.9453
23.3183 | -35.5114 | -34.9515
23.3244 | -35.5119 | -34.9576
23.3305 | -35.5124 | -34.9637
23.3366 | -35.513 | -34.9698
23.3427 | -35.5136 | -34.9762
23.3488 | -35.5142 | -34.9825
23.3549 | -35.5148 | -34.9889
23.361 | -35.5155 | -34.9946
23.3671 | -35.5162 | -35.001
23.3732 | -35.5168 | -35.0078
23.3793 | -35.5174 | -35.0145
23.3854 | -35.5181 | -35.0211
23.3915 | -35.5186 | -35.0278
23.3976 | -35.5193 | -35.0341
23.4037 | -35.5198 | -35.0399
23.4098 | -35.5204 | -35.0464
23.4159 | -35.5209 | -35.0522
23.422 | -35.5215 | -35.0579
23.4282 | -35.5221 | -35.0635
23.4343 | -35.5227 | -35.0701
23.4404 | -35.5233 | -35.0768
23.4465 | -35.5239 | -35.083
23.4526 | -35.5245 | -35.0892
23.4587 | -35.525 | -35.0958
23.4648 | -35.5256 | -35.1013
23.4709 | -35.5263 | -35.1075
23.477 | -35.5268 | -35.1141
23.4831 | -35.5274 | -35.1203
23.4892 | -35.5279 | -35.1266
23.4953 | -35.5285 | -35.1333
23.5014 | -35.529 | -35.1396
23.5075 | -35.5296 | -35.1457
23.5136 | -35.5302 | -35.1521
23.5197 | -35.5308 | -35.1581
23.5258 | -35.5314 | -35.1641
23.5319 | -35.532 | -35.1705
23.538 | -35.5325 | -35.1771
23.5441 | -35.5331 | -35.1837
23.5502 | -35.5338 | -35.1901
23.5563 | -35.5344 | -35.1965
23.5624 | -35.5349 | -35.2034
23.5686 | -35.5355 | -35.2093
23.5747 | -35.5361 | -35.216
23.5808 | -35.5368 | -35.222
23.5869 | -35.5374 | -35.2283
23.593 | -35.5379 | -35.2348
23.5991 | -35.5386 | -35.2414
23.6052 | -35.5391 | -35.2473
23.6113 | -35.5398 | -35.2533
23.6174 | -35.5403 | -35.2598
23.6235 | -35.5409 | -35.2657
23.6296 | -35.5415 | -35.2723
23.6357 | -35.5421 | -35.279
23.6418 | -35.5427 | -35.2849
23.6479 | -35.5433 | -35.2907
23.654 | -35.5439 | -35.2973
23.6601 | -35.5444 | -35.3037
23.6662 | -35.545 | -35.3098
23.6723 | -35.5455 | -35.3153
23.6784 | -35.5461 | -35.3215
23.6845 | -35.5467 | -35.3277
23.6906 | -35.5472 | -35.3342
23.6967 | -35.5479 | -35.3405
23.7028 | -35.5484 | -35.3469
23.7089 | -35.549 | -35.3526
23.7151 | -35.5496 | -35.3595
23.7212 | -35.5501 | -35.3658
23.7273 | -35.5507 | -35.3723
23.7334 | -35.5512 | -35.3787
23.7395 | -35.5519 | -35.385
23.7456 | -35.5525 | -35.3909
23.7517 | -35.5531 | -35.3977
23.7578 | -35.5536 | -35.4034
23.7639 | -35.5542 | -35.4097
23.77 | -35.5548 | -35.4159
23.7761 | -35.5553 | -35.4223
23.7822 | -35.5559 | -35.4283
23.7883 | -35.5565 | -35.4343
23.7944 | -35.5571 | -35.4405
23.8005 | -35.5576 | -35.4468
23.8066 | -35.5581 | -35.4533
23.8127 | -35.5588 | -35.4592
23.8188 | -35.5593 | -35.4658
23.8249 | -35.5598 | -35.4718
23.831 | -35.5603 | -35.4779
23.8371 | -35.561 | -35.4835
23.8432 | -35.5615 | -35.49
23.8493 | -35.5621 | -35.4959
23.8555 | -35.5627 | -35.5018
23.8616 | -35.5633 | -35.5083
23.8677 | -35.5639 | -35.5147
23.8738 | -35.5646 | -35.5206
23.8799 | -35.5651 | -35.5268
23.886 | -35.5656 | -35.5326
23.8921 | -35.5661 | -35.5389
23.8982 | -35.5667 | -35.5455
23.9043 | -35.5672 | -35.5517
23.9104 | -35.5678 | -35.5584
23.9165 | -35.5684 | -35.5647
23.9226 | -35.5689 | -35.57
23.9287 | -35.5694 | -35.5764
23.9348 | -35.57 | -35.5825
23.9409 | -35.5706 | -35.5882
23.947 | -35.5711 | -35.5942
23.9531 | -35.5717 | -35.5998
23.9592 | -35.5724 | -35.6066
23.9653 | -35.573 | -35.6128
23.9714 | -35.5735 | -35.6188
23.9775 | -35.5741 | -35.6247
23.9836 | -35.5746 | -35.6311
23.9897 | -35.5752 | -35.6371
23.9958 | -35.5758 | -35.6429
24.002 | -35.5763 | -35.649
24.0081 | -35.5769 | -35.6554
24.0142 | -35.5775 | -35.6615
24.0203 | -35.5781 | -35.6674
24.0264 | -35.5787 | -35.6732
24.0325 | -35.5793 | -35.6791
24.0386 | -35.5799 | -35.6854
24.0447 | -35.5804 | -35.6918
24.0508 | -35.581 | -35.6974
24.0569 | -35.5816 | -35.7037
24.063 | -35.5822 | -35.7095
24.0691 | -35.5828 | -35.7155
24.0752 | -35.5833 | -35.7219
24.0813 | -35.5839 | -35.7275
24.0874 | -35.5845 | -35.7339
24.0935 | -35.5851 | -35.7404
24.0996 | -35.5857 | -35.7466
24.1057 | -35.5862 | -35.753
24.1118 | -35.5868 | -35.7589
24.1179 | -35.5875 | -35.7654
24.124 | -35.588 | -35.7717
24.1301 | -35.5885 | -35.7774
24.1362 | -35.5891 | -35.7834
24.1424 | -35.5897 | -35.7897
24.1485 | -35.5903 | -35.7961
24.1546 | -35.5909 | -35.8023
24.1607 | -35.5914 | -35.8083
24.1668 | -35.5921 | -35.8143
24.1729 | -35.5927 | -35.8202
24.179 | -35.5933 | -35.8261
24.1851 | -35.5938 | -35.8325
24.1912 | -35.5943 | -35.8384
24.1973 | -35.5948 | -35.8447
24.2034 | -35.5951 | -35.8499
24.2095 | -35.5957 | -35.8562
24.2156 | -35.5962 | -35.8629
24.2217 | -35.5967 | -35.8687
24.2278 | -35.5973 | -35.8742
24.2339 | -35.5979 | -35.8806
24.24 | -35.5985 | -35.8872
24.2461 | -35.599 | -35.8931
24.2522 | -35.5996 | -35.8995
24.2583 | -35.6002 | -35.9052
24.2644 | -35.6008 | -35.9113
24.2705 | -35.6014 | -35.9169
24.2766 | -35.6019 | -35.9231
24.2827 | -35.6025 | -35.929
24.2889 | -35.6031 | -35.9355
24.295 | -35.6037 | -35.9418
24.3011 | -35.6042 | -35.9475
24.3072 | -35.6048 | -35.9531
24.3133 | -35.6053 | -35.959
24.3194 | -35.6059 | -35.9657
24.3255 | -35.6064 | -35.9719
24.3316 | -35.6069 | -35.9782
24.3377 | -35.6074 | -35.9845
24.3438 | -35.6079 | -35.9905
24.3499 | -35.6085 | -35.9966
24.356 | -35.609 | -36.0027
24.3621 | -35.6095 | -36.0091
24.3682 | -35.6101 | -36.0153
24.3743 | -35.6107 | -36.0212
24.3804 | -35.6112 | -36.0268
24.3865 | -35.6117 | -36.0328
24.3926 | -35.6123 | -36.0389
24.3987 | -35.6128 | -36.0454
24.4048 | -35.6133 | -36.0515
24.4109 | -35.6139 | -36.0568
24.417 | -35.6144 | -36.063
24.4231 | -35.615 | -36.0694
24.4293 | -35.6156 | -36.0755
24.4354 | -35.6162 | -36.0815
24.4415 | -35.6168 | -36.0873
24.4476 | -35.6174 | -36.0936
24.4537 | -35.618 | -36.099
24.4598 | -35.6186 | -36.105
24.4659 | -35.6191 | -36.1107
24.472 | -35.6197 | -36.1169
24.4781 | -35.6203 | -36.1229
24.4842 | -35.6207 | -36.1282
24.4903 | -35.6212 | -36.1339
24.4964 | -35.6218 | -36.1397
24.5025 | -35.6224 | -36.1456
24.5086 | -35.6229 | -36.152
24.5147 | -35.6235 | -36.1578
24.5208 | -35.624 | -36.1633
24.5269 | -35.6246 | -36.1694
24.533 | -35.6251 | -36.1752
24.5391 | -35.6257 | -36.1807
24.5452 | -35.6263 | -36.1864
24.5513 | -35.6268 | -36.1929
24.5574 | -35.6274 | -36.1989
24.5635 | -35.628 | -36.2049
24.5696 | -35.6286 | -36.2109
24.5758 | -35.6292 | -36.217
24.5819 | -35.6298 | -36.223
24.588 | -35.6304 | -36.2287
24.5941 | -35.6309 | -36.2344
24.6002 | -35.6315 | -36.2397
24.6063 | -35.632 | -36.2453
24.6124 | -35.6325 | -36.2508
24.6185 | -35.6331 | -36.257
24.6246 | -35.6337 | -36.2628
24.6307 | -35.6342 | -36.2687
24.6368 | -35.6347 | -36.275
24.6429 | -35.6352 | -36.2808
24.649 | -35.6358 | -36.2874
24.6551 | -35.6363 | -36.2931
24.6612 | -35.6367 | -36.2985
24.6673 | -35.6373 | -36.3045
24.6734 | -35.6378 | -36.3104
24.6795 | -35.6384 | -36.3167
24.6856 | -35.639 | -36.3231
24.6917 | -35.6396 | -36.3298
24.6978 | -35.6401 | -36.335
24.7039 | -35.6406 | -36.3407
24.71 | -35.6411 | -36.3472
24.7162 | -35.6417 | -36.3531
24.7223 | -35.6423 | -36.3591
24.7284 | -35.6428 | -36.3649
24.7345 | -35.6433 | -36.3705
24.7406 | -35.6438 | -36.3765
24.7467 | -35.6443 | -36.3829
24.7528 | -35.6449 | -36.3883
24.7589 | -35.6455 | -36.3944
24.765 | -35.646 | -36.3994
24.7711 | -35.6466 | -36.4053
24.7772 | -35.6472 | -36.4109
24.7833 | -35.6477 | -36.4168
24.7894 | -35.6484 | -36.4223
24.7955 | -35.6489 | -36.428
24.8016 | -35.6494 | -36.4333
24.8077 | -35.65 | -36.4394
24.8138 | -35.6505 | -36.4448
24.8199 | -35.6511 | -36.4507
24.826 | -35.6517 | -36.4568
24.8321 | -35.6523 | -36.4625
24.8382 | -35.6526 | -36.4679
24.8443 | -35.6532 | -36.4738
24.8504 | -35.6537 | -36.4797
24.8565 | -35.6543 | -36.4854
24.8627 | -35.6549 | -36.4909
24.8688 | -35.6553 | -36.4954
24.8749 | -35.6558 | -36.501
24.881 | -35.6563 | -36.507
24.8871 | -35.6569 | -36.5126
24.8932 | -35.6575 | -36.5186
24.8993 | -35.6579 | -36.5245
24.9054 | -35.6584 | -36.53
24.9115 | -35.6589 | -36.5357
24.9176 | -35.6594 | -36.5418
24.9237 | -35.66 | -36.5481
24.9298 | -35.6605 | -36.5542
24.9359 | -35.661 | -36.5607
24.942 | -35.6614 | -36.5666
24.9481 | -35.6619 | -36.5725
24.9542 | -35.6624 | -36.5788
24.9603 | -35.6629 | -36.5845
24.9664 | -35.6634 | -36.5904
24.9725 | -35.6639 | -36.5964
24.9786 | -35.6644 | -36.6014
24.9847 | -35.6649 | -36.6068
24.9908 | -35.6654 | -36.6126
24.9969 | -35.666 | -36.6185
25.0031 | -35.6666 | -36.6241
25.0092 | -35.6672 | -36.6293
25.0153 | -35.6677 | -36.6352
25.0214 | -35.6684 | -36.6409
25.0275 | -35.6689 | -36.647
25.0336 | -35.6695 | -36.653
25.0397 | -35.6701 | -36.6585
25.0458 | -35.6705 | -36.6642
25.0519 | -35.671 | -36.6697
25.058 | -35.6715 | -36.6754
25.0641 | -35.6721 | -36.6811
25.0702 | -35.6727 | -36.6864
25.0763 | -35.6733 | -36.6925
25.0824 | -35.6738 | -36.6984
25.0885 | -35.6744 | -36.7041
25.0946 | -35.6749 | -36.7099
25.1007 | -35.6753 | -36.7163
25.1068 | -35.6759 | -36.7222
25.1129 | -35.6764 | -36.7283
25.119 | -35.677 | -36.7335
25.1251 | -35.6775 | -36.7392
25.1312 | -35.6781 | -36.7454
25.1373 | -35.6787 | -36.7512
25.1435 | -35.6792 | -36.7572
25.1496 | -35.6797 | -36.7625
25.1557 | -35.6803 | -36.7687
25.1618 | -35.6808 | -36.7742
25.1679 | -35.6813 | -36.7798
25.174 | -35.6818 | -36.7851
25.1801 | -35.6824 | -36.7907
25.1862 | -35.6829 | -36.7965
25.1923 | -35.6835 | -36.8019
25.1984 | -35.684 | -36.8078
25.2045 | -35.6845 | -36.8138
25.2106 | -35.6851 | -36.8196
25.2167 | -35.6857 | -36.8258
25.2228 | -35.6862 | -36.8309
25.2289 | -35.6868 | -36.8366
25.235 | -35.6874 | -36.8419
25.2411 | -35.6879 | -36.8475
25.2472 | -35.6884 | -36.8534
25.2533 | -35.6889 | -36.8594
25.2594 | -35.6895 | -36.865
25.2655 | -35.69 | -36.8702
25.2716 | -35.6905 | -36.8756
25.2777 | -35.691 | -36.8809
25.2838 | -35.6915 | -36.8868
25.29 | -35.692 | -36.8921
25.2961 | -35.6925 | -36.8975
25.3022 | -35.693 | -36.9029
25.3083 | -35.6936 | -36.9086
25.3144 | -35.6942 | -36.9138
25.3205 | -35.6947 | -36.9199
25.3266 | -35.6952 | -36.9251
25.3327 | -35.6958 | -36.9313
25.3388 | -35.6962 | -36.9375
25.3449 | -35.6968 | -36.9428
25.351 | -35.6973 | -36.9486
25.3571 | -35.6977 | -36.9546
25.3632 | -35.6982 | -36.9597
25.3693 | -35.6988 | -36.9657
25.3754 | -35.6993 | -36.9714
25.3815 | -35.6998 | -36.9769
25.3876 | -35.7003 | -36.9826
25.3937 | -35.7008 | -36.9883
25.3998 | -35.7013 | -36.9937
25.4059 | -35.7018 | -36.9994
25.412 | -35.7024 | -37.0054
25.4181 | -35.7028 | -37.0106
25.4242 | -35.7033 | -37.016
25.4304 | -35.7038 | -37.021
25.4365 | -35.7043 | -37.0259
25.4426 | -35.7048 | -37.0315
25.4487 | -35.7053 | -37.0371
25.4548 | -35.7058 | -37.0424
25.4609 | -35.7062 | -37.0479
25.467 | -35.7068 | -37.0534
25.4731 | -35.7073 | -37.0586
25.4792 | -35.7078 | -37.0645
25.4853 | -35.7083 | -37.0702
25.4914 | -35.7089 | -37.0761
25.4975 | -35.7094 | -37.0816
25.5036 | -35.7098 | -37.0861
25.5097 | -35.7103 | -37.0914
25.5158 | -35.7108 | -37.0977
25.5219 | -35.7113 | -37.1033
25.528 | -35.7119 | -37.1087
25.5341 | -35.7124 | -37.1144
25.5402 | -35.7128 | -37.1204
25.5463 | -35.7134 | -37.1258
25.5524 | -35.714 | -37.1312
25.5585 | -35.7145 | -37.1371
25.5646 | -35.7149 | -37.1426
25.5707 | -35.7155 | -37.1482
25.5769 | -35.7159 | -37.1534
25.583 | -35.7165 | -37.1585
25.5891 | -35.7169 | -37.1645
25.5952 | -35.7175 | -37.1701
25.6013 | -35.718 | -37.1765
25.6074 | -35.7185 | -37.1816
25.6135 | -35.7191 | -37.1869
25.6196 | -35.7195 | -37.1926
25.6257 | -35.7201 | -37.1982
25.6318 | -35.7205 | -37.2038
25.6379 | -35.7209 | -37.2095
25.644 | -35.7214 | -37.2145
25.6501 | -35.722 | -37.2202
25.6562 | -35.7225 | -37.2265
25.6623 | -35.723 | -37.232
25.6684 | -35.7236 | -37.2375
25.6745 | -35.7242 | -37.243
25.6806 | -35.7248 | -37.2484
25.6867 | -35.7253 | -37.2538
25.6928 | -35.7257 | -37.2591
25.6989 | -35.7262 | -37.2651
25.705 | -35.7268 | -37.2709
25.7111 | -35.7274 | -37.2773
25.7173 | -35.7278 | -37.2833
25.7234 | -35.7283 | -37.2895
25.7295 | -35.7287 | -37.2948
25.7356 | -35.7292 | -37.3002
25.7417 | -35.7297 | -37.3058
25.7478 | -35.7302 | -37.3104
25.7539 | -35.7306 | -37.3163
25.76 | -35.7311 | -37.3215
25.7661 | -35.7315 | -37.3268
25.7722 | -35.732 | -37.332
25.7783 | -35.7326 | -37.3378
25.7844 | -35.7332 | -37.3432
25.7905 | -35.7338 | -37.349
25.7966 | -35.7342 | -37.3547
25.8027 | -35.7347 | -37.3603
25.8088 | -35.7351 | -37.365
25.8149 | -35.7357 | -37.3708
25.821 | -35.7362 | -37.3757
25.8271 | -35.7367 | -37.3815
25.8332 | -35.7371 | -37.3872
25.8393 | -35.7376 | -37.3935
25.8454 | -35.7381 | -37.3985
25.8515 | -35.7386 | -37.4044
25.8576 | -35.7392 | -37.4098
25.8638 | -35.7396 | -37.415
25.8699 | -35.7401 | -37.421
25.876 | -35.7405 | -37.4264
25.8821 | -35.741 | -37.4319
25.8882 | -35.7414 | -37.4375
25.8943 | -35.7418 | -37.4431
25.9004 | -35.7423 | -37.4484
25.9065 | -35.7428 | -37.4541
25.9126 | -35.7432 | -37.4599
25.9187 | -35.7438 | -37.4651
25.9248 | -35.7443 | -37.4702
25.9309 | -35.7448 | -37.4755
25.937 | -35.7453 | -37.4812
25.9431 | -35.7456 | -37.4864
25.9492 | -35.746 | -37.4914
25.9553 | -35.7465 | -37.4974
25.9614 | -35.7469 | -37.5025
25.9675 | -35.7474 | -37.5078
25.9736 | -35.7478 | -37.5139
25.9797 | -35.7483 | -37.5193
25.9858 | -35.7488 | -37.5242
25.9919 | -35.7493 | -37.5296
25.998 | -35.7499 | -37.5347
26.0042 | -35.7503 | -37.5402
26.0103 | -35.7507 | -37.5457
26.0164 | -35.7512 | -37.5513
26.0225 | -35.7517 | -37.5561
26.0286 | -35.7522 | -37.5613
26.0347 | -35.7527 | -37.5666
26.0408 | -35.7532 | -37.5721
26.0469 | -35.7537 | -37.5773
26.053 | -35.7542 | -37.583
26.0591 | -35.7547 | -37.5881
26.0652 | -35.7552 | -37.5939
26.0713 | -35.7556 | -37.5995
26.0774 | -35.7561 | -37.6051
26.0835 | -35.7567 | -37.6107
26.0896 | -35.7572 | -37.6167
26.0957 | -35.7578 | -37.6221
26.1018 | -35.7583 | -37.6274
26.1079 | -35.7587 | -37.6333
26.114 | -35.7593 | -37.6379
26.1201 | -35.7597 | -37.643
26.1262 | -35.7601 | -37.648
26.1323 | -35.7607 | -37.6538
26.1384 | -35.7611 | -37.6588
26.1445 | -35.7616 | -37.6644
26.1507 | -35.7621 | -37.6699
26.1568 | -35.7627 | -37.6757
26.1629 | -35.7631 | -37.6816
26.169 | -35.7636 | -37.6872
26.1751 | -35.764 | -37.6924
26.1812 | -35.7645 | -37.6977
26.1873 | -35.7651 | -37.7037
26.1934 | -35.7656 | -37.7095
26.1995 | -35.7661 | -37.7153
26.2056 | -35.7666 | -37.7206
26.2117 | -35.7671 | -37.7259
26.2178 | -35.7675 | -37.7308
26.2239 | -35.768 | -37.7364
26.23 | -35.7684 | -37.7422
26.2361 | -35.7689 | -37.7478
26.2422 | -35.7694 | -37.7535
26.2483 | -35.7699 | -37.7595
26.2544 | -35.7703 | -37.7649
26.2605 | -35.7708 | -37.7707
26.2666 | -35.7712 | -37.7758
26.2727 | -35.7718 | -37.7811
26.2788 | -35.7724 | -37.7862
26.2849 | -35.7729 | -37.7917
26.2911 | -35.7734 | -37.7966
26.2972 | -35.7739 | -37.8022
26.3033 | -35.7744 | -37.808
26.3094 | -35.7748 | -37.8135
26.3155 | -35.7752 | -37.8181
26.3216 | -35.7757 | -37.8239
26.3277 | -35.776 | -37.8294
26.3338 | -35.7765 | -37.8347
26.3399 | -35.777 | -37.8397
26.346 | -35.7775 | -37.8449
26.3521 | -35.7779 | -37.8507
26.3582 | -35.7784 | -37.8559
26.3643 | -35.7789 | -37.8613
26.3704 | -35.7795 | -37.8668
26.3765 | -35.7799 | -37.872
26.3826 | -35.7804 | -37.8772
26.3887 | -35.7809 | -37.8825
26.3948 | -35.7814 | -37.8889
26.4009 | -35.7818 | -37.8948
26.407 | -35.7824 | -37.9008
26.4131 | -35.7829 | -37.907
26.4192 | -35.7833 | -37.911
26.4253 | -35.7838 | -37.9169
26.4314 | -35.7842 | -37.9222
26.4376 | -35.7846 | -37.9284
26.4437 | -35.7851 | -37.9339
26.4498 | -35.7856 | -37.9394
26.4559 | -35.7861 | -37.9452
26.462 | -35.7866 | -37.9501
26.4681 | -35.7871 | -37.9554
26.4742 | -35.7875 | -37.9607
26.4803 | -35.788 | -37.9666
26.4864 | -35.7886 | -37.972
26.4925 | -35.789 | -37.9767
26.4986 | -35.7895 | -37.9816
26.5047 | -35.7898 | -37.9874
26.5108 | -35.7904 | -37.9937
26.5169 | -35.791 | -37.9993
26.523 | -35.7916 | -38.0042
26.5291 | -35.7922 | -38.0091
26.5352 | -35.7927 | -38.0148
26.5413 | -35.7931 | -38.0199
26.5474 | -35.7935 | -38.0245
26.5535 | -35.794 | -38.0306
26.5596 | -35.7945 | -38.0362
26.5657 | -35.795 | -38.0417
26.5718 | -35.7954 | -38.0471
26.578 | -35.7959 | -38.052
26.5841 | -35.7964 | -38.0567
26.5902 | -35.7969 | -38.062
26.5963 | -35.7973 | -38.0676
26.6024 | -35.7977 | -38.0724
26.6085 | -35.7982 | -38.0772
26.6146 | -35.7987 | -38.0824
26.6207 | -35.7991 | -38.0878
26.6268 | -35.7995 | -38.0933
26.6329 | -35.8001 | -38.0976
26.639 | -35.8005 | -38.1032
26.6451 | -35.801 | -38.109
26.6512 | -35.8015 | -38.1137
26.6573 | -35.802 | -38.1187
26.6634 | -35.8025 | -38.1233
26.6695 | -35.803 | -38.1287
26.6756 | -35.8034 | -38.134
26.6817 | -35.8038 | -38.1401
26.6878 | -35.8044 | -38.1453
26.6939 | -35.8049 | -38.1505
26.7 | -35.8053 | -38.1555
26.7061 | -35.8057 | -38.1605
26.7122 | -35.8061 | -38.1659
26.7183 | -35.8065 | -38.1713
26.7245 | -35.8071 | -38.1757
26.7306 | -35.8075 | -38.1813
26.7367 | -35.8079 | -38.1866
26.7428 | -35.8085 | -38.1922
26.7489 | -35.809 | -38.1982
26.755 | -35.8095 | -38.204
26.7611 | -35.81 | -38.2094
26.7672 | -35.8105 | -38.2153
26.7733 | -35.8109 | -38.2215
26.7794 | -35.8114 | -38.2269
26.7855 | -35.8118 | -38.232
26.7916 | -35.8121 | -38.2371
26.7977 | -35.8126 | -38.2428
26.8038 | -35.8131 | -38.2477
26.8099 | -35.8136 | -38.2528
26.816 | -35.814 | -38.2584
26.8221 | -35.8144 | -38.2635
26.8282 | -35.8148 | -38.2689
26.8343 | -35.8152 | -38.2746
26.8404 | -35.8156 | -38.2796
26.8465 | -35.816 | -38.2848
26.8526 | -35.8164 | -38.2898
26.8587 | -35.8169 | -38.2958
26.8649 | -35.8174 | -38.3004
26.871 | -35.8178 | -38.3052
26.8771 | -35.8181 | -38.309
26.8832 | -35.8186 | -38.3146
26.8893 | -35.8191 | -38.3201
26.8954 | -35.8196 | -38.3262
26.9015 | -35.8201 | -38.3311
26.9076 | -35.8204 | -38.3367
26.9137 | -35.821 | -38.3423
26.9198 | -35.8215 | -38.3477
26.9259 | -35.822 | -38.3532
26.932 | -35.8224 | -38.3588
26.9381 | -35.8229 | -38.3646
26.9442 | -35.8234 | -38.3692
26.9503 | -35.8238 | -38.3742
26.9564 | -35.8243 | -38.3789
26.9625 | -35.8248 | -38.3838
26.9686 | -35.8252 | -38.3893
26.9747 | -35.8257 | -38.3943
26.9808 | -35.8261 | -38.3991
26.9869 | -35.8265 | -38.4044
26.993 | -35.8269 | -38.41
26.9991 | -35.8273 | -38.416
27.0052 | -35.8278 | -38.4209
27.0114 | -35.8282 | -38.4261
27.0175 | -35.8287 | -38.4308
27.0236 | -35.8291 | -38.436
27.0297 | -35.8296 | -38.4417
27.0358 | -35.8301 | -38.4473
27.0419 | -35.8304 | -38.4526
27.048 | -35.8308 | -38.4575
27.0541 | -35.8313 | -38.463
27.0602 | -35.8318 | -38.4685
27.0663 | -35.8323 | -38.4739
27.0724 | -35.8327 | -38.4796
27.0785 | -35.8332 | -38.4851
27.0846 | -35.8335 | -38.4903
27.0907 | -35.8338 | -38.4954
27.0968 | -35.8342 | -38.5002
27.1029 | -35.8347 | -38.5051
27.109 | -35.8352 | -38.5104
27.1151 | -35.8355 | -38.515
27.1212 | -35.836 | -38.5206
27.1273 | -35.8364 | -38.5259
27.1334 | -35.8369 | -38.5315
27.1395 | -35.8374 | -38.5367
27.1456 | -35.8379 | -38.5428
27.1518 | -35.8383 | -38.5482
27.1579 | -35.8387 | -38.5528
27.164 | -35.8391 | -38.5581
27.1701 | -35.8394 | -38.5638
27.1762 | -35.8399 | -38.5691
27.1823 | -35.8404 | -38.5741
27.1884 | -35.8408 | -38.5782
27.1945 | -35.8412 | -38.5823
27.2006 | -35.8417 | -38.5869
27.2067 | -35.8422 | -38.5924
27.2128 | -35.8427 | -38.5969
27.2189 | -35.843 | -38.6018
27.225 | -35.8436 | -38.6071
27.2311 | -35.8442 | -38.6128
27.2372 | -35.8446 | -38.6181
27.2433 | -35.8451 | -38.6226
27.2494 | -35.8456 | -38.6279
27.2555 | -35.8461 | -38.6333
27.2616 | -35.8465 | -38.639
27.2677 | -35.8469 | -38.6444
27.2738 | -35.8474 | -38.6497
27.2799 | -35.8479 | -38.6545
27.286 | -35.8483 | -38.6593
27.2921 | -35.8488 | -38.6648
27.2983 | -35.8491 | -38.6692
27.3044 | -35.8496 | -38.6749
27.3105 | -35.85 | -38.6799
27.3166 | -35.8505 | -38.6846
27.3227 | -35.8509 | -38.6896
27.3288 | -35.8514 | -38.6944
27.3349 | -35.8519 | -38.6994
27.341 | -35.8523 | -38.7046
27.3471 | -35.8527 | -38.7096
27.3532 | -35.8531 | -38.7148
27.3593 | -35.8536 | -38.7205
27.3654 | -35.854 | -38.7255
27.3715 | -35.8543 | -38.7293
27.3776 | -35.8548 | -38.7346
27.3837 | -35.8552 | -38.7394
27.3898 | -35.8557 | -38.745
27.3959 | -35.8561 | -38.7502
27.402 | -35.8564 | -38.755
27.4081 | -35.8569 | -38.7609
27.4142 | -35.8573 | -38.7662
27.4203 | -35.8576 | -38.7715
27.4264 | -35.8581 | -38.7766
27.4325 | -35.8586 | -38.7815
27.4387 | -35.8591 | -38.786
27.4448 | -35.8596 | -38.791
27.4509 | -35.8601 | -38.796
27.457 | -35.8605 | -38.8014
27.4631 | -35.861 | -38.8068
27.4692 | -35.8614 | -38.8124
27.4753 | -35.8618 | -38.8169
27.4814 | -35.8621 | -38.8215
27.4875 | -35.8626 | -38.8262
27.4936 | -35.863 | -38.8323
27.4997 | -35.8634 | -38.837
27.5058 | -35.8639 | -38.8419
27.5119 | -35.8641 | -38.8481
27.518 | -35.8645 | -38.8534
27.5241 | -35.8649 | -38.8589
27.5302 | -35.8654 | -38.8647
27.5363 | -35.8658 | -38.8689
27.5424 | -35.8661 | -38.8741
27.5485 | -35.8666 | -38.8787
27.5546 | -35.867 | -38.8835
27.5607 | -35.8675 | -38.8883
27.5668 | -35.8678 | -38.893
27.5729 | -35.8683 | -38.8978
27.5791 | -35.8687 | -38.9033
27.5852 | -35.8691 | -38.9083
27.5913 | -35.8696 | -38.9131
27.5974 | -35.87 | -38.9189
27.6035 | -35.8704 | -38.9245
27.6096 | -35.8708 | -38.9296
27.6157 | -35.8712 | -38.9346
27.6218 | -35.8717 | -38.9393
27.6279 | -35.872 | -38.9454
27.634 | -35.8724 | -38.9504
27.6401 | -35.8727 | -38.956
27.6462 | -35.8732 | -38.9608
27.6523 | -35.8736 | -38.9667
27.6584 | -35.8739 | -38.9727
27.6645 | -35.8744 | -38.9784
27.6706 | -35.8747 | -38.9832
27.6767 | -35.8751 | -38.9891
27.6828 | -35.8755 | -38.9933
27.6889 | -35.8759 | -38.9982
27.695 | -35.8762 | -39.0043
27.7011 | -35.8766 | -39.0093
27.7072 | -35.877 | -39.0155
27.7133 | -35.8774 | -39.0209
27.7194 | -35.8776 | -39.0258
27.7256 | -35.8779 | -39.0304
27.7317 | -35.8783 | -39.0349
27.7378 | -35.8789 | -39.04
27.7439 | -35.8793 | -39.0458
27.75 | -35.8798 | -39.0506
27.7561 | -35.8802 | -39.0555
27.7622 | -35.8806 | -39.0603
27.7683 | -35.8811 | -39.0658
27.7744 | -35.8815 | -39.0708
27.7805 | -35.8819 | -39.0761
27.7866 | -35.8823 | -39.0812
27.7927 | -35.8828 | -39.0871
27.7988 | -35.8833 | -39.0917
27.8049 | -35.8838 | -39.0971
27.811 | -35.8842 | -39.1026
27.8171 | -35.8847 | -39.1079
27.8232 | -35.885 | -39.1129
27.8293 | -35.8854 | -39.1179
27.8354 | -35.8858 | -39.1234
27.8415 | -35.8862 | -39.129
27.8476 | -35.8867 | -39.1353
27.8537 | -35.8871 | -39.1413
27.8598 | -35.8876 | -39.1468
27.866 | -35.8881 | -39.1522
27.8721 | -35.8885 | -39.1569
27.8782 | -35.889 | -39.1614
27.8843 | -35.8895 | -39.1665
27.8904 | -35.8897 | -39.1713
27.8965 | -35.8902 | -39.1759
27.9026 | -35.8906 | -39.1816
27.9087 | -35.891 | -39.1867
27.9148 | -35.8914 | -39.1924
27.9209 | -35.8918 | -39.1977
27.927 | -35.8922 | -39.2024
27.9331 | -35.8925 | -39.2079
27.9392 | -35.8929 | -39.2133
27.9453 | -35.8934 | -39.218
27.9514 | -35.8936 | -39.2222
27.9575 | -35.894 | -39.2267
27.9636 | -35.8943 | -39.2322
27.9697 | -35.8948 | -39.2379
27.9758 | -35.8951 | -39.2429
27.9819 | -35.8954 | -39.2479
27.988 | -35.8959 | -39.2529
27.9941 | -35.8962 | -39.2578
28.0002 | -35.8965 | -39.262
28.0063 | -35.8969 | -39.2673
28.0125 | -35.8973 | -39.2723
28.0186 | -35.8976 | -39.2768
28.0247 | -35.898 | -39.2812
28.0308 | -35.8984 | -39.286
28.0369 | -35.8988 | -39.291
28.043 | -35.8991 | -39.2963
28.0491 | -35.8994 | -39.3014
28.0552 | -35.8997 | -39.3058
28.0613 | -35.9001 | -39.3102
28.0674 | -35.9006 | -39.3149
28.0735 | -35.9009 | -39.3207

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.6879 | -43.7503
21.4908 | -35.6198 | -36.0025
16.4948 | -34.9476 | -28.4201
12.6602 | -34.5533 | -21.9713
9.7171 | -34.3242 | -16.9296
7.4582 | -34.162 | -12.8412
5.7244 | -34.0496 | -9.5342
4.3936 | -33.9777 | -6.8634
3.3722 | -33.9633 | -4.7589
2.5883 | -33.9673 | -3.1661
1.9866 | -33.9581 | -1.8897
1.5248 | -33.962 | -0.86904
1.1703 | -33.9762 | -0.20381
1 | -33.9699 | 0.12775
0.89822 | -33.9614 | 0.4033
0.68941 | -33.9537 | 0.99202
0.52919 | -33.9718 | 1.4517
0.40612 | -33.9687 | 2.0574
0.31171 | -33.9536 | 2.5166
0.23928 | -33.9737 | 3.2395
0.18358 | -33.9873 | 3.6496
0.14091 | -33.983 | 4.5305
0.10819 | -34.0008 | 5.6786
0.083017 | -34.035 | 7.1483
0.063705 | -34.0866 | 9.0106
0.048923 | -34.1549 | 11.4907
0.037479 | -34.2575 | 14.6724
0.028801 | -34.4329 | 18.7124
0.022078 | -34.7449 | 24.115
0.017023 | -35.2354 | 30.5164
0.013018 | -36.0398 | 38.1756
0.010014 | -37.1417 | 46.0835

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003688 | -3.3086
0.007377 | -27.9295
0.011065 | -34.2222
0.014753 | -38.1508
0.018442 | -41.0096
0.02213 | -43.933
0.025818 | -45.7216
0.029507 | -47.3672
0.033195 | -48.814
0.036883 | -49.4903
0.040572 | -50.3593
0.04426 | -52.2977
0.047948 | -53.5355
0.051637 | -53.6599
0.055325 | -55.0639
0.059013 | -56.6758
0.062702 | -57.9609
0.06639 | -57.7625
0.070078 | -57.3897
0.073767 | -58.2257
0.077455 | -60.2207
0.081143 | -62.2803
0.084832 | -61.9467
0.08852 | -61.2507
0.092208 | -62.239
0.095897 | -62.8914
0.099585 | -62.4841
0.10327 | -62.3067
0.10696 | -62.5769
0.11065 | -62.9166
0.11434 | -64.532
0.11803 | -65.008
0.12172 | -64.4555
0.1254 | -64.7592
0.12909 | -65.1007
0.13278 | -65.2126
0.13647 | -64.7257
0.14016 | -65.5022
0.14385 | -65.6719
0.14753 | -66.4066
0.15122 | -66.7862
0.15491 | -66.7377
0.1586 | -67.1774
0.16229 | -67.0436
0.16598 | -67.4993
0.16966 | -68.4618
0.17335 | -68.6964
0.17704 | -69.9277
0.18073 | -70.0878
0.18442 | -69.4215
0.1881 | -69.3777
0.19179 | -70.2199
0.19548 | -70.613
0.19917 | -70.5558
0.20286 | -70.6429
0.20655 | -70.8203
0.21024 | -71.194
0.21392 | -71.3986
0.21761 | -72.0163
0.2213 | -71.483
0.22499 | -71.6964
0.22868 | -71.8372
0.23236 | -71.4882
0.23605 | -73.093
0.23974 | -72.9525
0.24343 | -72.0128
0.24712 | -72.1818
0.25081 | -72.6872
0.2545 | -72.868
0.25818 | -71.7789
0.26187 | -71.76
0.26556 | -73.2013
0.26925 | -73.9413
0.27294 | -74.5457
0.27663 | -74.6729
0.28031 | -74.007
0.284 | -73.983
0.28769 | -74.3026
0.29138 | -74.5265
0.29507 | -73.6791
0.29875 | -74.1178
0.30244 | -74.6022
0.30613 | -74.765
0.30982 | -74.9819
0.31351 | -74.4572
0.3172 | -74.8487
0.32088 | -74.8756
0.32457 | -75.2066
0.32826 | -75.2423
0.33195 | -74.3501
0.33564 | -74.6794
0.33933 | -75.9143
0.34302 | -76.7263
0.3467 | -76.9411
0.35039 | -76.1487
0.35408 | -74.9165
0.35777 | -74.8026
0.36146 | -76.0218
0.36514 | -76.8623
0.36883 | -77.3272
0.37252 | -76.0407
0.37621 | -76.0364
0.3799 | -76.5558
0.38359 | -76.1485
0.38727 | -76.0067
0.39096 | -76.2629
0.39465 | -77.132
0.39834 | -77.7455
0.40203 | -78.3501
0.40572 | -77.1048
0.40941 | -76.566
0.41309 | -77.1662
0.41678 | -76.808
0.42047 | -77.1566
0.42416 | -77.5962
0.42785 | -77.0858
0.43154 | -78.0962
0.43522 | -78.4796
0.43891 | -77.5696
0.4426 | -76.8196
0.44629 | -76.9411
0.44998 | -78.3294
0.45366 | -78.4254
0.45735 | -78.2083
0.46104 | -79.0266
0.46473 | -78.659
0.46842 | -78.4664
0.47211 | -79.073
0.4758 | -79.9466
0.47948 | -79.1694
0.48317 | -79.8251
0.48686 | -80.1714
0.49055 | -79.3286
0.49424 | -78.6272
0.49793 | -79.1305
0.50161 | -79.115
0.5053 | -78.3377
0.50899 | -78.8189
0.51268 | -79.5496
0.51637 | -80.0072
0.52006 | -80.3999
0.52374 | -80.2866
0.52743 | -80.0591
0.53112 | -79.7395
0.53481 | -79.895
0.5385 | -80.0775
0.54219 | -80.4455
0.54587 | -80.7387
0.54956 | -80.6949
0.55325 | -80.5455
0.55694 | -80.52
0.56063 | -80.8461
0.56432 | -81.0587
0.568 | -81.0272
0.57169 | -80.9678
0.57538 | -80.7426
0.57907 | -80.5942
0.58276 | -80.6813
0.58644 | -80.8141
0.59013 | -80.8277
0.59382 | -80.6463
0.59751 | -80.7401
0.6012 | -81.1026
0.60489 | -81.5088
0.60857 | -81.759
0.61226 | -81.8586
0.61595 | -81.8302
0.61964 | -81.9247
0.62333 | -81.9552
0.62702 | -81.6301
0.6307 | -81.2327
0.63439 | -81.0126
0.63808 | -81.0751
0.64177 | -81.3207
0.64546 | -81.6282
0.64915 | -81.4357
0.65284 | -81.2531
0.65652 | -81.5911
0.66021 | -82.0185
0.6639 | -82.0263
0.66759 | -82.066
0.67128 | -82.3878
0.67497 | -82.7795
0.67865 | -82.7459
0.68234 | -82.8054
0.68603 | -83.1847
0.68972 | -83.3367
0.69341 | -83.162
0.6971 | -83.087
0.70078 | -83.1325
0.70447 | -82.8267
0.70816 | -82.2649
0.71185 | -81.8995
0.71554 | -81.8886
0.71923 | -82.2108
0.72291 | -82.378
0.7266 | -82.3031
0.73029 | -82.1124
0.73398 | -82.2305
0.73767 | -83.1931
0.74135 | -84.1477
0.74504 | -84.1167
0.74873 | -83.7715
0.75242 | -83.5533
0.75611 | -83.5343
0.7598 | -83.661
0.76348 | -83.5305
0.76717 | -83.0844
0.77086 | -83.0321
0.77455 | -83.5829
0.77824 | -83.9377
0.78193 | -83.8276
0.78561 | -83.8774
0.7893 | -84.0773
0.79299 | -83.7742
0.79668 | -83.0308
0.80037 | -83.07
0.80406 | -83.6719
0.80775 | -83.9695
0.81143 | -83.954
0.81512 | -83.943
0.81881 | -83.7512
0.8225 | -84.088
0.82619 | -84.9627
0.82988 | -85.2689
0.83356 | -85.1068
0.83725 | -85.1857
0.84094 | -85.4717
0.84463 | -85.4143
0.84832 | -85.0124
0.85201 | -84.5843
0.85569 | -84.3395
0.85938 | -84.4503
0.86307 | -84.7757
0.86676 | -84.59
0.87045 | -84.4709
0.87414 | -84.912
0.87782 | -84.9228
0.88151 | -84.5802
0.8852 | -84.4633
0.88889 | -84.7545
0.89258 | -85.0278
0.89627 | -85.1411
0.89995 | -85.2129
0.90364 | -85.3298
0.90733 | -85.9559
0.91102 | -86.0618
0.91471 | -85.7967
0.9184 | -85.35
0.92208 | -84.7621
0.92577 | -84.3387
0.92946 | -84.0225
0.93315 | -84.3397
0.93684 | -85.0669
0.94053 | -85.2375
0.94421 | -85.1268
0.9479 | -85.0641
0.95159 | -85.0405
0.95528 | -85.4258
0.95897 | -86.0211
0.96266 | -86.0214
0.96634 | -85.643
0.97003 | -85.7299
0.97372 | -85.8081
0.97741 | -85.2932
0.9811 | -85.0988
0.98479 | -85.6911
0.98847 | -86.4653
0.99216 | -86.3398
0.99585 | -85.841
0.99954 | -85.7142
1.0032 | -85.9243
1.0069 | -85.9979
1.0106 | -85.9735
1.0143 | -86.2001
1.018 | -86.7004
1.0217 | -86.6238
1.0254 | -86.0048
1.029 | -85.632
1.0327 | -86.0352
1.0364 | -86.663
1.0401 | -86.7064
1.0438 | -86.2412
1.0475 | -85.8253
1.0512 | -85.9537
1.0549 | -86.3787
1.0586 | -86.7773
1.0622 | -86.688
1.0659 | -86.5532
1.0696 | -86.5064
1.0733 | -86.7269
1.077 | -87.1193
1.0807 | -87.421
1.0844 | -87.4423
1.0881 | -87.1709
1.0917 | -86.7859
1.0954 | -86.6179
1.0991 | -86.8861
1.1028 | -87.4838
1.1065 | -87.9561
1.1102 | -87.8888
1.1139 | -87.3936
1.1176 | -86.7901
1.1213 | -86.2871
1.1249 | -86.0557
1.1286 | -86.2259
1.1323 | -86.6537
1.136 | -86.9234
1.1397 | -86.9025
1.1434 | -86.876
1.1471 | -87.0553
1.1508 | -87.3799
1.1544 | -87.4425
1.1581 | -87.0319
1.1618 | -86.5885
1.1655 | -86.4567
1.1692 | -86.6751
1.1729 | -87.2531
1.1766 | -87.9412
1.1803 | -88.1768
1.184 | -87.8973
1.1876 | -87.6024
1.1913 | -87.6959
1.195 | -88.1666
1.1987 | -88.5841
1.2024 | -88.6612
1.2061 | -88.5295
1.2098 | -88.3591
1.2135 | -88.3378
1.2172 | -88.4763
1.2208 | -88.5061
1.2245 | -88.3217
1.2282 | -88.0439
1.2319 | -87.7302
1.2356 | -87.5612
1.2393 | -87.7135
1.243 | -87.9801
1.2467 | -87.9886
1.2503 | -87.8111
1.254 | -87.76
1.2577 | -87.8261
1.2614 | -87.8538
1.2651 | -87.8755
1.2688 | -87.928
1.2725 | -88.017
1.2762 | -88.174
1.2799 | -88.2916
1.2835 | -88.2524
1.2872 | -88.1217
1.2909 | -88.0233
1.2946 | -88.0164
1.2983 | -88.0741
1.302 | -88.1723
1.3057 | -88.2826
1.3094 | -88.3012
1.313 | -88.3109
1.3167 | -88.525
1.3204 | -88.8985
1.3241 | -89.2029
1.3278 | -89.3692
1.3315 | -89.4786
1.3352 | -89.5803
1.3389 | -89.5264
1.3426 | -89.1524
1.3462 | -88.6751
1.3499 | -88.4604
1.3536 | -88.6812
1.3573 | -89.1974
1.361 | -89.4723
1.3647 | -89.0896
1.3684 | -88.3882
1.3721 | -87.9236
1.3757 | -87.8424
1.3794 | -88.0348
1.3831 | -88.4794
1.3868 | -89.0565
1.3905 | -89.4028
1.3942 | -89.3545
1.3979 | -89.1996
1.4016 | -89.228
1.4053 | -89.379
1.4089 | -89.4035
1.4126 | -89.2225
1.4163 | -89.0142
1.42 | -89.0353
1.4237 | -89.292
1.4274 | -89.5269
1.4311 | -89.5769
1.4348 | -89.5066
1.4385 | -89.4135
1.4421 | -89.2882
1.4458 | -89.0982
1.4495 | -88.9288
1.4532 | -88.907
1.4569 | -89.0257
1.4606 | -89.1567
1.4643 | -89.2513
1.468 | -89.4326
1.4716 | -89.7342
1.4753 | -90.0456
1.479 | -90.3031
1.4827 | -90.4817
1.4864 | -90.5512
1.4901 | -90.5902
1.4938 | -90.6044
1.4975 | -90.4711
1.5012 | -90.2064
1.5048 | -90.0027
1.5085 | -89.9816
1.5122 | -90.0354
1.5159 | -90.0251
1.5196 | -89.9139
1.5233 | -89.692
1.527 | -89.5129
1.5307 | -89.5227
1.5343 | -89.6132
1.538 | -89.6308
1.5417 | -89.6037
1.5454 | -89.6854
1.5491 | -89.9055
1.5528 | -90.1171
1.5565 | -90.2568
1.5602 | -90.343
1.5639 | -90.3105
1.5675 | -90.209
1.5712 | -90.1948
1.5749 | -90.2817
1.5786 | -90.3052
1.5823 | -90.1274
1.586 | -89.9971
1.5897 | -90.1924
1.5934 | -90.3778
1.597 | -90.4125
1.6007 | -90.3207
1.6044 | -90.1969
1.6081 | -90.1111
1.6118 | -90.1085
1.6155 | -90.1822
1.6192 | -90.2571
1.6229 | -90.2385
1.6266 | -90.1176
1.6302 | -89.9546
1.6339 | -89.8182
1.6376 | -89.7758
1.6413 | -89.8209
1.645 | -89.8773
1.6487 | -89.8831
1.6524 | -89.8513
1.6561 | -89.802
1.6598 | -89.7723
1.6634 | -89.8121
1.6671 | -89.8799
1.6708 | -89.8943
1.6745 | -89.9052
1.6782 | -90.0385
1.6819 | -90.3402
1.6856 | -90.724
1.6893 | -91.0486
1.6929 | -91.2063
1.6966 | -91.2014
1.7003 | -91.1303
1.704 | -91.0371
1.7077 | -90.9451
1.7114 | -90.8889
1.7151 | -90.8673
1.7188 | -90.8596
1.7225 | -90.833
1.7261 | -90.7574
1.7298 | -90.6496
1.7335 | -90.5552
1.7372 | -90.5371
1.7409 | -90.5783
1.7446 | -90.6439
1.7483 | -90.7402
1.752 | -90.8394
1.7556 | -90.9183
1.7593 | -90.9621
1.763 | -91.0077
1.7667 | -91.0834
1.7704 | -91.1351
1.7741 | -91.1212
1.7778 | -91.0648
1.7815 | -91.0108
1.7852 | -90.9657
1.7888 | -90.8629
1.7925 | -90.6851
1.7962 | -90.5056
1.7999 | -90.4499
1.8036 | -90.607
1.8073 | -90.9458
1.811 | -91.312
1.8147 | -91.5288
1.8183 | -91.51
1.822 | -91.3788
1.8257 | -91.2838
1.8294 | -91.2795
1.8331 | -91.3803
1.8368 | -91.593
1.8405 | -91.8744
1.8442 | -92.1087
1.8479 | -92.1815
1.8515 | -92.0744
1.8552 | -91.8103
1.8589 | -91.5485
1.8626 | -91.437
1.8663 | -91.508
1.87 | -91.7002
1.8737 | -91.9259
1.8774 | -92.1232
1.8811 | -92.2582
1.8847 | -92.3441
1.8884 | -92.4302
1.8921 | -92.5166
1.8958 | -92.578
1.8995 | -92.5602
1.9032 | -92.4478
1.9069 | -92.2694
1.9106 | -92.0765
1.9142 | -91.908
1.9179 | -91.7643
1.9216 | -91.6588
1.9253 | -91.622
1.929 | -91.676
1.9327 | -91.8197
1.9364 | -92.0416
1.9401 | -92.307
1.9438 | -92.5054
1.9474 | -92.5155
1.9511 | -92.3154
1.9548 | -91.9921
1.9585 | -91.6853
1.9622 | -91.5078
1.9659 | -91.4715
1.9696 | -91.5359
1.9733 | -91.7052
1.9769 | -91.9737
1.9806 | -92.2262
1.9843 | -92.3149
1.988 | -92.2146
1.9917 | -92.0341
1.9954 | -91.8457
1.9991 | -91.7049
2.0028 | -91.6474
2.0065 | -91.7051
2.0101 | -91.8557
2.0138 | -92.0161
2.0175 | -92.098
2.0212 | -92.0787
2.0249 | -92.0305
2.0286 | -92.0485
2.0323 | -92.1445
2.036 | -92.2842
2.0396 | -92.3997
2.0433 | -92.3906
2.047 | -92.2282
2.0507 | -92.0486
2.0544 | -91.9996
2.0581 | -92.0906
2.0618 | -92.2562
2.0655 | -92.4462
2.0692 | -92.6198
2.0728 | -92.7056
2.0765 | -92.6649
2.0802 | -92.5647
2.0839 | -92.5257
2.0876 | -92.6288
2.0913 | -92.8564
2.095 | -93.119
2.0987 | -93.2212
2.1024 | -93.03
2.106 | -92.6556
2.1097 | -92.3224
2.1134 | -92.1406
2.1171 | -92.1041
2.1208 | -92.156
2.1245 | -92.2663
2.1282 | -92.4044
2.1319 | -92.5162
2.1355 | -92.5978
2.1392 | -92.6413
2.1429 | -92.6646
2.1466 | -92.6832
2.1503 | -92.7029
2.154 | -92.7235
2.1577 | -92.7529
2.1614 | -92.7809
2.1651 | -92.8014
2.1687 | -92.8189
2.1724 | -92.8472
2.1761 | -92.9028
2.1798 | -93.0137
2.1835 | -93.1816
2.1872 | -93.3671
2.1909 | -93.5024
2.1946 | -93.5562
2.1982 | -93.5417
2.2019 | -93.4962
2.2056 | -93.4472
2.2093 | -93.4299
2.213 | -93.45
2.2167 | -93.4681
2.2204 | -93.4268
2.2241 | -93.3108
2.2278 | -93.1433
2.2314 | -92.9519
2.2351 | -92.7563
2.2388 | -92.6236
2.2425 | -92.6211
2.2462 | -92.743
2.2499 | -92.9116
2.2536 | -93.0278
2.2573 | -93.0345
2.2609 | -92.9669
2.2646 | -92.8816
2.2683 | -92.8372
2.272 | -92.8488
2.2757 | -92.9032
2.2794 | -92.9733
2.2831 | -93.0299
2.2868 | -93.0619
2.2905 | -93.0505
2.2941 | -92.9788
2.2978 | -92.8549
2.3015 | -92.7084
2.3052 | -92.5982
2.3089 | -92.5864
2.3126 | -92.6983
2.3163 | -92.9218
2.32 | -93.2016
2.3237 | -93.4622
2.3273 | -93.6555
2.331 | -93.7289
2.3347 | -93.6606
2.3384 | -93.4966
2.3421 | -93.3359
2.3458 | -93.2519
2.3495 | -93.2475
2.3532 | -93.2863
2.3568 | -93.3295
2.3605 | -93.3549
2.3642 | -93.3651
2.3679 | -93.3242
2.3716 | -93.2478
2.3753 | -93.1836
2.379 | -93.18
2.3827 | -93.2517
2.3864 | -93.3752
2.39 | -93.4883
2.3937 | -93.5408
2.3974 | -93.5309
2.4011 | -93.5164
2.4048 | -93.5252
2.4085 | -93.5571
2.4122 | -93.5598
2.4159 | -93.5089
2.4195 | -93.4278
2.4232 | -93.3811
2.4269 | -93.4019
2.4306 | -93.4699
2.4343 | -93.5457
2.438 | -93.5939
2.4417 | -93.624
2.4454 | -93.6735
2.4491 | -93.7217
2.4527 | -93.7288
2.4564 | -93.7005
2.4601 | -93.6902
2.4638 | -93.7536
2.4675 | -93.9117
2.4712 | -94.1384
2.4749 | -94.3784
2.4786 | -94.5622
2.4823 | -94.6081
2.4859 | -94.4723
2.4896 | -94.234
2.4933 | -94.0068
2.497 | -93.8773
2.5007 | -93.8813
2.5044 | -93.9649
2.5081 | -94.0587
2.5118 | -94.1084
2.5154 | -94.072
2.5191 | -93.9418
2.5228 | -93.7153
2.5265 | -93.4337
2.5302 | -93.1836
2.5339 | -93.06
2.5376 | -93.1016
2.5413 | -93.2882
2.545 | -93.5518
2.5486 | -93.799
2.5523 | -93.9998
2.556 | -94.1773
2.5597 | -94.3499
2.5634 | -94.5135
2.5671 | -94.6194
2.5708 | -94.6358
2.5745 | -94.5629
2.5781 | -94.4307
2.5818 | -94.2772
2.5855 | -94.1582
2.5892 | -94.1228
2.5929 | -94.1776
2.5966 | -94.3053
2.6003 | -94.4679
2.604 | -94.6323
2.6077 | -94.7677
2.6113 | -94.8384
2.615 | -94.8409
2.6187 | -94.7779
2.6224 | -94.6745
2.6261 | -94.5589
2.6298 | -94.4692
2.6335 | -94.4114
2.6372 | -94.3882
2.6408 | -94.4195
2.6445 | -94.5142
2.6482 | -94.651
2.6519 | -94.7837
2.6556 | -94.8612
2.6593 | -94.8847
2.663 | -94.9239
2.6667 | -94.9412
2.6704 | -94.9301
2.674 | -94.893
2.6777 | -94.8333
2.6814 | -94.7399
2.6851 | -94.6231
2.6888 | -94.5083
2.6925 | -94.422
2.6962 | -94.3895
2.6999 | -94.4205
2.7035 | -94.4912
2.7072 | -94.5696
2.7109 | -94.6396
2.7146 | -94.6872
2.7183 | -94.7198
2.722 | -94.7358
2.7257 | -94.7555
2.7294 | -94.7827
2.7331 | -94.7972
2.7367 | -94.7994
2.7404 | -94.8069
2.7441 | -94.8209
2.7478 | -94.839
2.7515 | -94.8579
2.7552 | -94.8784
2.7589 | -94.8921
2.7626 | -94.8999
2.7663 | -94.8901
2.7699 | -94.864
2.7736 | -94.8038
2.7773 | -94.6943
2.781 | -94.5572
2.7847 | -94.4025
2.7884 | -94.245
2.7921 | -94.1168
2.7958 | -94.0474
2.7994 | -94.0546
2.8031 | -94.1376
2.8068 | -94.2928
2.8105 | -94.4752
2.8142 | -94.6234
2.8179 | -94.7079
2.8216 | -94.7206
2.8253 | -94.6924
2.829 | -94.6596
2.8326 | -94.6711
2.8363 | -94.7411
2.84 | -94.8501
2.8437 | -94.9677
2.8474 | -95.0616
2.8511 | -95.1244
2.8548 | -95.1696
2.8585 | -95.2299
2.8621 | -95.3267
2.8658 | -95.4316
2.8695 | -95.5261
2.8732 | -95.5976
2.8769 | -95.6354
2.8806 | -95.6285
2.8843 | -95.5897
2.888 | -95.5306
2.8917 | -95.4739
2.8953 | -95.4477
2.899 | -95.4608
2.9027 | -95.5167
2.9064 | -95.6204
2.9101 | -95.7565
2.9138 | -95.8986
2.9175 | -96.0066
2.9212 | -96.0487
2.9249 | -96.0265
2.9285 | -95.9595
2.9322 | -95.8744
2.9359 | -95.7815
2.9396 | -95.6913
2.9433 | -95.6081
2.947 | -95.5416
2.9507 | -95.4694
2.9544 | -95.3851
2.958 | -95.2785
2.9617 | -95.1612
2.9654 | -95.0534
2.9691 | -94.9613
2.9728 | -94.9077
2.9765 | -94.8908
2.9802 | -94.8928
2.9839 | -94.9225
2.9876 | -94.9536
2.9912 | -94.9771
2.9949 | -95.0064
2.9986 | -95.0753
3.0023 | -95.1938
3.006 | -95.3508
3.0097 | -95.5148
3.0134 | -95.6496
3.0171 | -95.7539
3.0207 | -95.8052
3.0244 | -95.8133
3.0281 | -95.7861
3.0318 | -95.7267
3.0355 | -95.6817
3.0392 | -95.6736
3.0429 | -95.6809
3.0466 | -95.6803
3.0503 | -95.6452
3.0539 | -95.58
3.0576 | -95.5142
3.0613 | -95.4795
3.065 | -95.4926
3.0687 | -95.5748
3.0724 | -95.7197
3.0761 | -95.9133
3.0798 | -96.1187
3.0834 | -96.2827
3.0871 | -96.3529
3.0908 | -96.3195
3.0945 | -96.1975
3.0982 | -96.023
3.1019 | -95.8429
3.1056 | -95.6959
3.1093 | -95.614
3.113 | -95.6052
3.1166 | -95.6595
3.1203 | -95.734
3.124 | -95.7795
3.1277 | -95.7699
3.1314 | -95.7025
3.1351 | -95.6007
3.1388 | -95.5106
3.1425 | -95.4656
3.1462 | -95.4833
3.1498 | -95.5411
3.1535 | -95.5968
3.1572 | -95.6165
3.1609 | -95.5608
3.1646 | -95.4163
3.1683 | -95.2287
3.172 | -95.078
3.1757 | -95.0156
3.1793 | -95.0612
3.183 | -95.2223
3.1867 | -95.4444
3.1904 | -95.6258
3.1941 | -95.7884
3.1978 | -95.8941
3.2015 | -95.9337
3.2052 | -95.9165
3.2089 | -95.8336
3.2125 | -95.7065
3.2162 | -95.5599
3.2199 | -95.4173
3.2236 | -95.3004
3.2273 | -95.2526
3.231 | -95.2968
3.2347 | -95.4178
3.2384 | -95.6076
3.242 | -95.8299
3.2457 | -96.0499
3.2494 | -96.2291
3.2531 | -96.3608
3.2568 | -96.4307
3.2605 | -96.4528
3.2642 | -96.4608
3.2679 | -96.4615
3.2716 | -96.4607
3.2752 | -96.4624
3.2789 | -96.4468
3.2826 | -96.3894
3.2863 | -96.283
3.29 | -96.1491
3.2937 | -95.9903
3.2974 | -95.8158
3.3011 | -95.6566
3.3047 | -95.534
3.3084 | -95.4499
3.3121 | -95.4207
3.3158 | -95.4453
3.3195 | -95.5159
3.3232 | -95.6286
3.3269 | -95.7648
3.3306 | -95.9173
3.3343 | -96.0787
3.3379 | -96.2319
3.3416 | -96.3539
3.3453 | -96.4215
3.349 | -96.4484
3.3527 | -96.4317
3.3564 | -96.376
3.3601 | -96.3068
3.3638 | -96.2425
3.3674 | -96.1931
3.3711 | -96.1844
3.3748 | -96.2243
3.3785 | -96.3133
3.3822 | -96.4413
3.3859 | -96.5763
3.3896 | -96.7009
3.3933 | -96.7993
3.397 | -96.8587
3.4006 | -96.902
3.4043 | -96.9381
3.408 | -96.984
3.4117 | -97.0251
3.4154 | -97.0642
3.4191 | -97.0973
3.4228 | -97.113
3.4265 | -97.1125
3.4302 | -97.107
3.4338 | -97.0842
3.4375 | -97.0372
3.4412 | -96.9737
3.4449 | -96.9039
3.4486 | -96.8403
3.4523 | -96.7904
3.456 | -96.7353
3.4597 | -96.6775
3.4633 | -96.6142
3.467 | -96.5583
3.4707 | -96.5292
3.4744 | -96.5452
3.4781 | -96.6186
3.4818 | -96.7352
3.4855 | -96.8628
3.4892 | -96.9765
3.4929 | -97.0458
3.4965 | -97.059
3.5002 | -97.0071
3.5039 | -96.9008
3.5076 | -96.777
3.5113 | -96.6749
3.515 | -96.6204
3.5187 | -96.6275
3.5224 | -96.6726
3.526 | -96.7346
3.5297 | -96.7861
3.5334 | -96.8146
3.5371 | -96.8305
3.5408 | -96.8386
3.5445 | -96.85
3.5482 | -96.8839
3.5519 | -96.9254
3.5556 | -96.9696
3.5592 | -97.0081
3.5629 | -97.0108
3.5666 | -96.9727
3.5703 | -96.9162
3.574 | -96.8519
3.5777 | -96.7845
3.5814 | -96.7289
3.5851 | -96.6866
3.5888 | -96.6487
3.5924 | -96.6088
3.5961 | -96.5773
3.5998 | -96.5663
3.6035 | -96.582
3.6072 | -96.6364
3.6109 | -96.7148
3.6146 | -96.7879
3.6183 | -96.8626
3.6219 | -96.9285
3.6256 | -96.9819
3.6293 | -97.0463
3.633 | -97.132
3.6367 | -97.2081
3.6404 | -97.2881
3.6441 | -97.3625
3.6478 | -97.4211
3.6515 | -97.4565
3.6551 | -97.4381
3.6588 | -97.364
3.6625 | -97.2401
3.6662 | -97.1102
3.6699 | -97.001
3.6736 | -96.9255
3.6773 | -96.9072
3.681 | -96.9427
3.6846 | -97.0153
3.6883 | -97.1031
3.692 | -97.1938
3.6957 | -97.2555
3.6994 | -97.2912
3.7031 | -97.3076
3.7068 | -97.3026
3.7105 | -97.3004
3.7142 | -97.2942
3.7178 | -97.2984
3.7215 | -97.3062
3.7252 | -97.3084
3.7289 | -97.3171
3.7326 | -97.3392
3.7363 | -97.3753
3.74 | -97.4208
3.7437 | -97.4796
3.7473 | -97.5533
3.751 | -97.6255
3.7547 | -97.6825
3.7584 | -97.7167
3.7621 | -97.73
3.7658 | -97.7116
3.7695 | -97.6496
3.7732 | -97.5597
3.7769 | -97.4501
3.7805 | -97.3577
3.7842 | -97.3029
3.7879 | -97.2882
3.7916 | -97.314
3.7953 | -97.3684
3.799 | -97.4407
3.8027 | -97.5225
3.8064 | -97.5945
3.8101 | -97.6454
3.8137 | -97.6747
3.8174 | -97.6767
3.8211 | -97.6387
3.8248 | -97.5608
3.8285 | -97.4653
3.8322 | -97.3655
3.8359 | -97.2822
3.8396 | -97.225
3.8432 | -97.2121
3.8469 | -97.2623
3.8506 | -97.3687
3.8543 | -97.5142
3.858 | -97.6737
3.8617 | -97.8207
3.8654 | -97.9399
3.8691 | -98.0253
3.8728 | -98.0794
3.8764 | -98.1059
3.8801 | -98.1107
3.8838 | -98.097
3.8875 | -98.0613
3.8912 | -98.0078
3.8949 | -97.9434
3.8986 | -97.8671
3.9023 | -97.7819
3.9059 | -97.6884
3.9096 | -97.6035
3.9133 | -97.5411
3.917 | -97.5033
3.9207 | -97.4923
3.9244 | -97.5101
3.9281 | -97.5411
3.9318 | -97.5848
3.9355 | -97.643
3.9391 | -97.7166
3.9428 | -97.7971
3.9465 | -97.8753
3.9502 | -97.9575
3.9539 | -98.0274
3.9576 | -98.0709
3.9613 | -98.0835
3.965 | -98.0733
3.9686 | -98.0433
3.9723 | -97.9934
3.976 | -97.9327
3.9797 | -97.8704
3.9834 | -97.814
3.9871 | -97.7798
3.9908 | -97.7844
3.9945 | -97.819
3.9982 | -97.8806
4.0018 | -97.9559
4.0055 | -98.039
4.0092 | -98.1258
4.0129 | -98.2023
4.0166 | -98.2672
4.0203 | -98.3142
4.024 | -98.3384
4.0277 | -98.342
4.0314 | -98.3239
4.035 | -98.297
4.0387 | -98.2735
4.0424 | -98.2637
4.0461 | -98.2675
4.0498 | -98.2803
4.0535 | -98.3069
4.0572 | -98.3405
4.0609 | -98.3671
4.0645 | -98.388
4.0682 | -98.4
4.0719 | -98.4034
4.0756 | -98.3967
4.0793 | -98.3904
4.083 | -98.3881
4.0867 | -98.387
4.0904 | -98.3794
4.0941 | -98.38
4.0977 | -98.3945
4.1014 | -98.4275
4.1051 | -98.4678
4.1088 | -98.508
4.1125 | -98.5464
4.1162 | -98.5682
4.1199 | -98.5772
4.1236 | -98.5722
4.1272 | -98.5429
4.1309 | -98.4995
4.1346 | -98.4424
4.1383 | -98.3823
4.142 | -98.3093
4.1457 | -98.2414
4.1494 | -98.184
4.1531 | -98.1551
4.1568 | -98.1518
4.1604 | -98.1793
4.1641 | -98.2407
4.1678 | -98.3402
4.1715 | -98.468
4.1752 | -98.6
4.1789 | -98.7243
4.1826 | -98.8204
4.1863 | -98.8738
4.1899 | -98.8824
4.1936 | -98.8529
4.1973 | -98.7841
4.201 | -98.6894
4.2047 | -98.579
4.2084 | -98.4646
4.2121 | -98.3683
4.2158 | -98.2963
4.2195 | -98.2568
4.2231 | -98.2492
4.2268 | -98.2757
4.2305 | -98.3299
4.2342 | -98.4033
4.2379 | -98.4988
4.2416 | -98.5965
4.2453 | -98.6782
4.249 | -98.7295
4.2527 | -98.7434
4.2563 | -98.6948
4.26 | -98.6627
4.2637 | -98.6163
4.2674 | -98.5562
4.2711 | -98.4976
4.2748 | -98.4332
4.2785 | -98.3696
4.2822 | -98.3086
4.2858 | -98.2557
4.2895 | -98.226
4.2932 | -98.2114
4.2969 | -98.2097
4.3006 | -98.2149
4.3043 | -98.2204
4.308 | -98.2324
4.3117 | -98.2503
4.3154 | -98.2796
4.319 | -98.3202
4.3227 | -98.3788
4.3264 | -98.4584
4.3301 | -98.5527
4.3338 | -98.6562
4.3375 | -98.7573
4.3412 | -98.841
4.3449 | -98.9181
4.3485 | -98.9903
4.3522 | -99.0518
4.3559 | -99.0926
4.3596 | -99.1065
4.3633 | -99.1036
4.367 | -99.089
4.3707 | -99.0672
4.3744 | -99.0406
4.3781 | -99.0099
4.3817 | -98.9846
4.3854 | -98.9678
4.3891 | -98.9595
4.3928 | -98.9481
4.3965 | -98.936
4.4002 | -98.9223
4.4039 | -98.9069
4.4076 | -98.8851
4.4112 | -98.8534
4.4149 | -98.8034
4.4186 | -98.7461
4.4223 | -98.6884
4.426 | -98.6285
4.4297 | -98.5729
4.4334 | -98.5334
4.4371 | -98.5022
4.4408 | -98.4939
4.4444 | -98.5058
4.4481 | -98.5312
4.4518 | -98.5519
4.4555 | -98.5593
4.4592 | -98.5546
4.4629 | -98.5384
4.4666 | -98.5169
4.4703 | -98.4916
4.474 | -98.4668
4.4776 | -98.4572
4.4813 | -98.4689
4.485 | -98.494
4.4887 | -98.516
4.4924 | -98.5404
4.4961 | -98.5598
4.4998 | -98.5612
4.5035 | -98.5436
4.5071 | -98.5152
4.5108 | -98.4697
4.5145 | -98.4244
4.5182 | -98.3935
4.5219 | -98.3776
4.5256 | -98.3703
4.5293 | -98.3677
4.533 | -98.3808
4.5367 | -98.4033
4.5403 | -98.4398
4.544 | -98.483
4.5477 | -98.5251
4.5514 | -98.563
4.5551 | -98.5947
4.5588 | -98.626
4.5625 | -98.652
4.5662 | -98.6849
4.5698 | -98.7234
4.5735 | -98.7712
4.5772 | -98.8261
4.5809 | -98.8858
4.5846 | -98.9371
4.5883 | -98.968
4.592 | -98.9834
4.5957 | -98.9865
4.5994 | -98.9794
4.603 | -98.9588
4.6067 | -98.936
4.6104 | -98.9231
4.6141 | -98.9297
4.6178 | -98.9581
4.6215 | -99.0171
4.6252 | -99.1126
4.6289 | -99.231
4.6326 | -99.3552
4.6362 | -99.4751
4.6399 | -99.5854
4.6436 | -99.6566
4.6473 | -99.6875
4.651 | -99.6682
4.6547 | -99.6004
4.6584 | -99.5019
4.6621 | -99.3854
4.6657 | -99.2639
4.6694 | -99.1418
4.6731 | -99.0316
4.6768 | -98.9453
4.6805 | -98.8784
4.6842 | -98.8308
4.6879 | -98.7988
4.6916 | -98.7895
4.6953 | -98.8018
4.6989 | -98.8303
4.7026 | -98.8849
4.7063 | -98.9464
4.71 | -98.9998
4.7137 | -99.0471
4.7174 | -99.0834
4.7211 | -99.1132
4.7248 | -99.1476
4.7284 | -99.1799
4.7321 | -99.2037
4.7358 | -99.2108
4.7395 | -99.2037
4.7432 | -99.172
4.7469 | -99.1098
4.7506 | -99.0444
4.7543 | -98.9869
4.758 | -98.9542
4.7616 | -98.9442
4.7653 | -98.9559
4.769 | -98.9899
4.7727 | -99.0287
4.7764 | -99.0752
4.7801 | -99.1211
4.7838 | -99.1564
4.7875 | -99.1276
4.7911 | -99.1192
4.7948 | -99.0908
4.7985 | -99.0466
4.8022 | -98.9945
4.8059 | -98.9413
4.8096 | -98.8976
4.8133 | -98.8624
4.817 | -98.8378
4.8207 | -98.8307
4.8243 | -98.8301
4.828 | -98.8358
4.8317 | -98.8492
4.8354 | -98.8632
4.8391 | -98.8879
4.8428 | -98.9178
4.8465 | -98.9489
4.8502 | -98.9778
4.8539 | -99.0046
4.8575 | -99.0266
4.8612 | -99.055
4.8649 | -99.0877
4.8686 | -99.119
4.8723 | -99.1491
4.876 | -99.1759
4.8797 | -99.1969
4.8834 | -99.2055
4.887 | -99.2025
4.8907 | -99.1994
4.8944 | -99.1992
4.8981 | -99.2073
4.9018 | -99.2349
4.9055 | -99.2903
4.9092 | -99.3678
4.9129 | -99.4694
4.9166 | -99.577
4.9202 | -99.6801
4.9239 | -99.773
4.9276 | -99.8485
4.9313 | -99.8972
4.935 | -99.9114
4.9387 | -99.8925
4.9424 | -99.854
4.9461 | -99.8096
4.9497 | -99.7606
4.9534 | -99.7076
4.9571 | -99.6627
4.9608 | -99.6358
4.9645 | -99.6227
4.9682 | -99.6113
4.9719 | -99.604
4.9756 | -99.5991
4.9793 | -99.5904
4.9829 | -99.5747
4.9866 | -99.5563
4.9903 | -99.533
4.994 | -99.5068
4.9977 | -99.4812
5.0014 | -99.4503
5.0051 | -99.424
5.0088 | -99.4119
5.0124 | -99.4174
5.0161 | -99.4246
5.0198 | -99.4359
5.0235 | -99.4462
5.0272 | -99.4519
5.0309 | -99.4418
5.0346 | -99.4087
5.0383 | -99.3524
5.042 | -99.2817
5.0456 | -99.2057
5.0493 | -99.133
5.053 | -99.0666
5.0567 | -99.0226
5.0604 | -98.9993
5.0641 | -98.9985
5.0678 | -99.0226
5.0715 | -99.0641
5.0751 | -99.1245
5.0788 | -99.1921
5.0825 | -99.2562
5.0862 | -99.3067
5.0899 | -99.3445
5.0936 | -99.3798
5.0973 | -99.3986
5.101 | -99.4033
5.1047 | -99.4064
5.1083 | -99.4108
5.112 | -99.4145
5.1157 | -99.4345
5.1194 | -99.4675
5.1231 | -99.5099
5.1268 | -99.5547
5.1305 | -99.6069
5.1342 | -99.6572
5.1379 | -99.7081
5.1415 | -99.7562
5.1452 | -99.798
5.1489 | -99.8329
5.1526 | -99.8509
5.1563 | -99.862
5.16 | -99.8641
5.1637 | -99.8581
5.1674 | -99.8435
5.171 | -99.8184
5.1747 | -99.7968
5.1784 | -99.7732
5.1821 | -99.7489
5.1858 | -99.7306
5.1895 | -99.7244
5.1932 | -99.7236
5.1969 | -99.7321
5.2006 | -99.7533
5.2042 | -99.7963
5.2079 | -99.8478
5.2116 | -99.9021
5.2153 | -99.9637
5.219 | -100.0245
5.2227 | -100.0792
5.2264 | -100.1271
5.2301 | -100.1608
5.2337 | -100.1694
5.2374 | -100.1574
5.2411 | -100.132
5.2448 | -100.0967
5.2485 | -100.0532
5.2522 | -100.0054
5.2559 | -99.9704
5.2596 | -99.9482
5.2633 | -99.9377
5.2669 | -99.937
5.2706 | -99.953
5.2743 | -99.9753
5.278 | -100.0051
5.2817 | -100.0344
5.2854 | -100.0714
5.2891 | -100.1179
5.2928 | -100.161
5.2965 | -100.2053
5.3001 | -100.2489
5.3038 | -100.2955
5.3075 | -100.3437
5.3112 | -100.3901
5.3149 | -100.4392
5.3186 | -100.4757
5.3223 | -100.509
5.326 | -100.5369
5.3296 | -100.5627
5.3333 | -100.583
5.337 | -100.5866
5.3407 | -100.578
5.3444 | -100.5647
5.3481 | -100.5533
5.3518 | -100.5286
5.3555 | -100.5045
5.3592 | -100.4694
5.3628 | -100.4305
5.3665 | -100.389
5.3702 | -100.3439
5.3739 | -100.3032
5.3776 | -100.2631
5.3813 | -100.2375
5.385 | -100.2218
5.3887 | -100.2176
5.3923 | -100.2315
5.396 | -100.2504
5.3997 | -100.2796
5.4034 | -100.3155
5.4071 | -100.3469
5.4108 | -100.369
5.4145 | -100.3837
5.4182 | -100.3891
5.4219 | -100.3799
5.4255 | -100.3574
5.4292 | -100.3154
5.4329 | -100.259
5.4366 | -100.2011
5.4403 | -100.1444
5.444 | -100.0908
5.4477 | -100.0343
5.4514 | -99.9884
5.455 | -99.9546
5.4587 | -99.9313
5.4624 | -99.9204
5.4661 | -99.9177
5.4698 | -99.9219
5.4735 | -99.9369
5.4772 | -99.953
5.4809 | -99.9719
5.4846 | -99.9955
5.4882 | -100.0174
5.4919 | -100.0464
5.4956 | -100.0796
5.4993 | -100.1029
5.503 | -100.125
5.5067 | -100.1506
5.5104 | -100.1739
5.5141 | -100.1986
5.5178 | -100.2192
5.5214 | -100.2373
5.5251 | -100.2563
5.5288 | -100.2759
5.5325 | -100.3
5.5362 | -100.3216
5.5399 | -100.3426
5.5436 | -100.3594
5.5473 | -100.3751
5.5509 | -100.3833
5.5546 | -100.3964
5.5583 | -100.4056
5.562 | -100.4105
5.5657 | -100.4211
5.5694 | -100.4279
5.5731 | -100.4378
5.5768 | -100.4319
5.5805 | -100.419
5.5841 | -100.3988
5.5878 | -100.3648
5.5915 | -100.3229
5.5952 | -100.2765
5.5989 | -100.2355
5.6026 | -100.2048
5.6063 | -100.1844
5.61 | -100.1759
5.6136 | -100.1853
5.6173 | -100.2041
5.621 | -100.2336
5.6247 | -100.2732
5.6284 | -100.3212
5.6321 | -100.3705
5.6358 | -100.422
5.6395 | -100.4649
5.6432 | -100.5086
5.6468 | -100.5497
5.6505 | -100.5768
5.6542 | -100.6011
5.6579 | -100.618
5.6616 | -100.6347
5.6653 | -100.6456
5.669 | -100.6563
5.6727 | -100.6679
5.6763 | -100.6789
5.68 | -100.6961
5.6837 | -100.7127
5.6874 | -100.7333
5.6911 | -100.7635
5.6948 | -100.7974
5.6985 | -100.8306
5.7022 | -100.8615
5.7059 | -100.8786
5.7095 | -100.8877
5.7132 | -100.8878
5.7169 | -100.8732
5.7206 | -100.8486
5.7243 | -100.8134
5.728 | -100.7604
5.7317 | -100.7074
5.7354 | -100.6422
5.739 | -100.5654
5.7427 | -100.4933
5.7464 | -100.4245
5.7501 | -100.358
5.7538 | -100.2964
5.7575 | -100.2341
5.7612 | -100.1785
5.7649 | -100.1318
5.7686 | -100.0981
5.7722 | -100.0766
5.7759 | -100.0727
5.7796 | -100.0888
5.7833 | -100.1235
5.787 | -100.1778
5.7907 | -100.2476
5.7944 | -100.3204
5.7981 | -100.3912
5.8018 | -100.4632
5.8054 | -100.5288
5.8091 | -100.5866
5.8128 | -100.6327
5.8165 | -100.6543
5.8202 | -100.6701
5.8239 | -100.6771
5.8276 | -100.6661
5.8313 | -100.6406
5.8349 | -100.6086
5.8386 | -100.5692
5.8423 | -100.5248
5.846 | -100.4726
5.8497 | -100.4157
5.8534 | -100.3673
5.8571 | -100.3269
5.8608 | -100.2963
5.8645 | -100.2805
5.8681 | -100.2759
5.8718 | -100.2831
5.8755 | -100.3054
5.8792 | -100.3376
5.8829 | -100.3736
5.8866 | -100.418
5.8903 | -100.4664
5.894 | -100.5087
5.8976 | -100.5442
5.9013 | -100.5719
5.905 | -100.5901
5.9087 | -100.6018
5.9124 | -100.6035
5.9161 | -100.5961
5.9198 | -100.5899
5.9235 | -100.5792
5.9272 | -100.5717
5.9308 | -100.5715
5.9345 | -100.5827
5.9382 | -100.6059
5.9419 | -100.6413
5.9456 | -100.6816
5.9493 | -100.7247
5.953 | -100.763
5.9567 | -100.8006
5.9604 | -100.8424
5.964 | -100.8794
5.9677 | -100.9012
5.9714 | -100.9119
5.9751 | -100.9169
5.9788 | -100.9145
5.9825 | -100.9029
5.9862 | -100.8941
5.9899 | -100.8894
5.9935 | -100.8938
5.9972 | -100.9054
6.0009 | -100.9163
6.0046 | -100.9279
6.0083 | -100.9461
6.012 | -100.9591
6.0157 | -100.9727
6.0194 | -100.9739
6.0231 | -100.9607
6.0267 | -100.9368
6.0304 | -100.8986
6.0341 | -100.8431
6.0378 | -100.7838
6.0415 | -100.7261
6.0452 | -100.6774
6.0489 | -100.6265
6.0526 | -100.5846
6.0562 | -100.558
6.0599 | -100.5462
6.0636 | -100.5493
6.0673 | -100.5592
6.071 | -100.5813
6.0747 | -100.6199
6.0784 | -100.6703
6.0821 | -100.7277
6.0858 | -100.7931
6.0894 | -100.8636
6.0931 | -100.9285
6.0968 | -100.991
6.1005 | -101.0464
6.1042 | -101.0856
6.1079 | -101.1079
6.1116 | -101.115
6.1153 | -101.1011
6.1189 | -101.0669
6.1226 | -101.0299
6.1263 | -100.9914
6.13 | -100.9515
6.1337 | -100.918
6.1374 | -100.8968
6.1411 | -100.8992
6.1448 | -100.919
6.1485 | -100.9558
6.1521 | -101.0141
6.1558 | -101.0878
6.1595 | -101.1774
6.1632 | -101.2676
6.1669 | -101.3543
6.1706 | -101.434
6.1743 | -101.5031
6.178 | -101.5573
6.1817 | -101.6012
6.1853 | -101.632
6.189 | -101.644
6.1927 | -101.6442
6.1964 | -101.6416
6.2001 | -101.6318
6.2038 | -101.6183
6.2075 | -101.6064
6.2112 | -101.5918
6.2148 | -101.5747
6.2185 | -101.5598
6.2222 | -101.5406
6.2259 | -101.5187
6.2296 | -101.4954
6.2333 | -101.4673
6.237 | -101.4434
6.2407 | -101.4213
6.2444 | -101.3921
6.248 | -101.3626
6.2517 | -101.3297
6.2554 | -101.2945
6.2591 | -101.2688
6.2628 | -101.2537
6.2665 | -101.2464
6.2702 | -101.2448
6.2739 | -101.2564
6.2775 | -101.2838
6.2812 | -101.3156
6.2849 | -101.3411
6.2886 | -101.3611
6.2923 | -101.3705
6.296 | -101.3724
6.2997 | -101.367
6.3034 | -101.3582
6.3071 | -101.3439
6.3107 | -101.3375
6.3144 | -101.3425
6.3181 | -101.3604
6.3218 | -101.3911
6.3255 | -101.4304
6.3292 | -101.4741
6.3329 | -101.5226
6.3366 | -101.5707
6.3402 | -101.6149
6.3439 | -101.6509
6.3476 | -101.6817
6.3513 | -101.6935
6.355 | -101.7001
6.3587 | -101.7057
6.3624 | -101.7136
6.3661 | -101.7156
6.3698 | -101.719
6.3734 | -101.7267
6.3771 | -101.7403
6.3808 | -101.7582
6.3845 | -101.7715
6.3882 | -101.7838
6.3919 | -101.788
6.3956 | -101.7814
6.3993 | -101.7683
6.403 | -101.7467
6.4066 | -101.7141
6.4103 | -101.6771
6.414 | -101.6315
6.4177 | -101.5832
6.4214 | -101.5453
6.4251 | -101.514
6.4288 | -101.4933
6.4325 | -101.4787
6.4361 | -101.4722
6.4398 | -101.4839
6.4435 | -101.5017
6.4472 | -101.5282
6.4509 | -101.5534
6.4546 | -101.5833
6.4583 | -101.6088
6.462 | -101.6317
6.4657 | -101.6477
6.4693 | -101.653
6.473 | -101.6487
6.4767 | -101.6357
6.4804 | -101.6129
6.4841 | -101.5809
6.4878 | -101.5454
6.4915 | -101.5157
6.4952 | -101.4858
6.4988 | -101.4604
6.5025 | -101.4425
6.5062 | -101.4316
6.5099 | -101.4239
6.5136 | -101.4233
6.5173 | -101.4266
6.521 | -101.4269
6.5247 | -101.4297
6.5284 | -101.4343
6.532 | -101.4445
6.5357 | -101.4551
6.5394 | -101.4613
6.5431 | -101.4731
6.5468 | -101.4906
6.5505 | -101.5054
6.5542 | -101.5166
6.5579 | -101.5292
6.5615 | -101.5391
6.5652 | -101.5518
6.5689 | -101.567
6.5726 | -101.581
6.5763 | -101.5969
6.58 | -101.6198
6.5837 | -101.6443
6.5874 | -101.6646
6.5911 | -101.6822
6.5947 | -101.7104
6.5984 | -101.7334
6.6021 | -101.7596
6.6058 | -101.7758
6.6095 | -101.7817
6.6132 | -101.7861
6.6169 | -101.7818
6.6206 | -101.7641
6.6243 | -101.7359
6.6279 | -101.7029
6.6316 | -101.6731
6.6353 | -101.6398
6.639 | -101.6052
6.6427 | -101.5745
6.6464 | -101.5497
6.6501 | -101.5277
6.6538 | -101.5051
6.6574 | -101.4825
6.6611 | -101.4655
6.6648 | -101.4503
6.6685 | -101.4292
6.6722 | -101.4116
6.6759 | -101.4002
6.6796 | -101.3934
6.6833 | -101.3911
6.687 | -101.3967
6.6906 | -101.4118
6.6943 | -101.4332
6.698 | -101.4672
6.7017 | -101.5053
6.7054 | -101.5519
6.7091 | -101.6099
6.7128 | -101.6696
6.7165 | -101.7314
6.7201 | -101.7962
6.7238 | -101.8573
6.7275 | -101.9084
6.7312 | -101.9554
6.7349 | -101.9959
6.7386 | -102.023
6.7423 | -102.0419
6.746 | -102.0491
6.7497 | -102.0491
6.7533 | -102.0385
6.757 | -102.0234
6.7607 | -101.9991
6.7644 | -101.9717
6.7681 | -101.9494
6.7718 | -101.9234
6.7755 | -101.8959
6.7792 | -101.8706
6.7828 | -101.8556
6.7865 | -101.8467
6.7902 | -101.8373
6.7939 | -101.8328
6.7976 | -101.8325
6.8013 | -101.8426
6.805 | -101.8626
6.8087 | -101.8803
6.8124 | -101.8946
6.816 | -101.9119
6.8197 | -101.9262
6.8234 | -101.9249
6.8271 | -101.9192
6.8308 | -101.9022
6.8345 | -101.8833
6.8382 | -101.8642
6.8419 | -101.8372
6.8456 | -101.8108
6.8492 | -101.7883
6.8529 | -101.7607
6.8566 | -101.7292
6.8603 | -101.7035
6.864 | -101.6832
6.8677 | -101.6688
6.8714 | -101.6557
6.8751 | -101.6482
6.8787 | -101.6465
6.8824 | -101.6468
6.8861 | -101.6537
6.8898 | -101.6624
6.8935 | -101.6729
6.8972 | -101.6883
6.9009 | -101.6991
6.9046 | -101.7017
6.9083 | -101.703
6.9119 | -101.7068
6.9156 | -101.6967
6.9193 | -101.6911
6.923 | -101.6846
6.9267 | -101.6863
6.9304 | -101.6856
6.9341 | -101.6834
6.9378 | -101.685
6.9414 | -101.6927
6.9451 | -101.7084
6.9488 | -101.7317
6.9525 | -101.7611
6.9562 | -101.806
6.9599 | -101.8561
6.9636 | -101.9075
6.9673 | -101.9628
6.971 | -102.0227
6.9746 | -102.0848
6.9783 | -102.1396
6.982 | -102.1908
6.9857 | -102.2355
6.9894 | -102.2663
6.9931 | -102.2894
6.9968 | -102.2986
7.0005 | -102.2972
7.0041 | -102.2881
7.0078 | -102.2726
7.0115 | -102.2474
7.0152 | -102.2077
7.0189 | -102.1695
7.0226 | -102.1303
7.0263 | -102.0862
7.03 | -102.0408
7.0337 | -101.9963
7.0373 | -101.9502
7.041 | -101.9125
7.0447 | -101.8745
7.0484 | -101.8357
7.0521 | -101.7979
7.0558 | -101.7693
7.0595 | -101.7453
7.0632 | -101.7213
7.0669 | -101.7109
7.0705 | -101.7121
7.0742 | -101.7263
7.0779 | -101.7495
7.0816 | -101.7859
7.0853 | -101.8332
7.089 | -101.8801
7.0927 | -101.9302
7.0964 | -101.9854
7.1 | -102.0375
7.1037 | -102.0857
7.1074 | -102.1283
7.1111 | -102.161
7.1148 | -102.1911
7.1185 | -102.2161
7.1222 | -102.2372
7.1259 | -102.2505
7.1296 | -102.2562
7.1332 | -102.2594
7.1369 | -102.2543
7.1406 | -102.2393
7.1443 | -102.2202
7.148 | -102.1964
7.1517 | -102.1745
7.1554 | -102.1535
7.1591 | -102.1329
7.1627 | -102.1196
7.1664 | -102.1223
7.1701 | -102.1376
7.1738 | -102.1591
7.1775 | -102.1857
7.1812 | -102.2233
7.1849 | -102.2729
7.1886 | -102.3227
7.1923 | -102.3693
7.1959 | -102.418
7.1996 | -102.4595
7.2033 | -102.4873
7.207 | -102.5062
7.2107 | -102.5157
7.2144 | -102.508
7.2181 | -102.4923
7.2218 | -102.465
7.2255 | -102.4321
7.2291 | -102.3933
7.2328 | -102.3489
7.2365 | -102.3056
7.2402 | -102.268
7.2439 | -102.2332
7.2476 | -102.2025
7.2513 | -102.1794
7.255 | -102.1593
7.2586 | -102.1455
7.2623 | -102.14
7.266 | -102.1424
7.2697 | -102.1394
7.2734 | -102.1373
7.2771 | -102.1487
7.2808 | -102.1649
7.2845 | -102.1909
7.2882 | -102.2148
7.2918 | -102.236
7.2955 | -102.2596
7.2992 | -102.2861
7.3029 | -102.3099
7.3066 | -102.3278
7.3103 | -102.3437
7.314 | -102.3555
7.3177 | -102.3599
7.3213 | -102.3625
7.325 | -102.3569
7.3287 | -102.3473
7.3324 | -102.3347
7.3361 | -102.3262
7.3398 | -102.3221
7.3435 | -102.3157
7.3472 | -102.3134
7.3509 | -102.3205
7.3545 | -102.336
7.3582 | -102.3542
7.3619 | -102.3802
7.3656 | -102.4088
7.3693 | -102.4371
7.373 | -102.4626
7.3767 | -102.4905
7.3804 | -102.5126
7.384 | -102.5246
7.3877 | -102.5324
7.3914 | -102.5345
7.3951 | -102.5284
7.3988 | -102.5133
7.4025 | -102.4936
7.4062 | -102.4626
7.4099 | -102.4236
7.4136 | -102.3875
7.4172 | -102.35
7.4209 | -102.3134
7.4246 | -102.2829
7.4283 | -102.259
7.432 | -102.2471
7.4357 | -102.2433
7.4394 | -102.2449
7.4431 | -102.2514
7.4467 | -102.2712
7.4504 | -102.2805
7.4541 | -102.2988
7.4578 | -102.3175
7.4615 | -102.3381
7.4652 | -102.3569
7.4689 | -102.3783
7.4726 | -102.398
7.4763 | -102.4186
7.4799 | -102.4421
7.4836 | -102.4693
7.4873 | -102.4967
7.491 | -102.527
7.4947 | -102.5595
7.4984 | -102.5959
7.5021 | -102.6345
7.5058 | -102.6757
7.5095 | -102.7193
7.5131 | -102.762
7.5168 | -102.8036
7.5205 | -102.8437
7.5242 | -102.8756
7.5279 | -102.9007
7.5316 | -102.9233
7.5353 | -102.9388
7.539 | -102.9451
7.5426 | -102.9377
7.5463 | -102.9206
7.55 | -102.8947
7.5537 | -102.8649
7.5574 | -102.8329
7.5611 | -102.793
7.5648 | -102.7495
7.5685 | -102.6927
7.5722 | -102.6318
7.5758 | -102.5699
7.5795 | -102.5149
7.5832 | -102.4601
7.5869 | -102.4104
7.5906 | -102.3636
7.5943 | -102.3189
7.598 | -102.2747
7.6017 | -102.2409
7.6053 | -102.2162
7.609 | -102.1963
7.6127 | -102.186
7.6164 | -102.1835
7.6201 | -102.1884
7.6238 | -102.193
7.6275 | -102.2012
7.6312 | -102.2127
7.6349 | -102.2281
7.6385 | -102.2453
7.6422 | -102.26
7.6459 | -102.2708
7.6496 | -102.2815
7.6533 | -102.293
7.657 | -102.299
7.6607 | -102.2981
7.6644 | -102.2934
7.6681 | -102.2837
7.6717 | -102.2721
7.6754 | -102.2621
7.6791 | -102.2514
7.6828 | -102.2467
7.6865 | -102.2408
7.6902 | -102.2372
7.6939 | -102.2321
7.6976 | -102.232
7.7012 | -102.2357
7.7049 | -102.2368
7.7086 | -102.2385
7.7123 | -102.2414
7.716 | -102.2403
7.7197 | -102.2366
7.7234 | -102.2368
7.7271 | -102.2314
7.7308 | -102.2241
7.7344 | -102.2216
7.7381 | -102.2252
7.7418 | -102.2324
7.7455 | -102.243
7.7492 | -102.2544
7.7529 | -102.2698
7.7566 | -102.2959
7.7603 | -102.3251
7.7639 | -102.3586
7.7676 | -102.396
7.7713 | -102.4431
7.775 | -102.4892
7.7787 | -102.5353
7.7824 | -102.5793
7.7861 | -102.6213
7.7898 | -102.6628
7.7935 | -102.7063
7.7971 | -102.7438
7.8008 | -102.7791
7.8045 | -102.8088
7.8082 | -102.8335
7.8119 | -102.8512
7.8156 | -102.863
7.8193 | -102.8709
7.823 | -102.8758
7.8266 | -102.8778
7.8303 | -102.8749
7.834 | -102.8653
7.8377 | -102.8482
7.8414 | -102.8338
7.8451 | -102.8198
7.8488 | -102.7968
7.8525 | -102.7734
7.8562 | -102.7446
7.8598 | -102.7214
7.8635 | -102.698
7.8672 | -102.6766
7.8709 | -102.6533
7.8746 | -102.6337
7.8783 | -102.6203
7.882 | -102.6112
7.8857 | -102.605
7.8894 | -102.6065
7.893 | -102.6156
7.8967 | -102.6303
7.9004 | -102.65
7.9041 | -102.6765
7.9078 | -102.7047
7.9115 | -102.7398
7.9152 | -102.7776
7.9189 | -102.8144
7.9225 | -102.8512
7.9262 | -102.8906
7.9299 | -102.9305
7.9336 | -102.9718
7.9373 | -103
7.941 | -103.018
7.9447 | -103.0304
7.9484 | -103.0391
7.9521 | -103.0351
7.9557 | -103.0183
7.9594 | -102.9884
7.9631 | -102.9524
7.9668 | -102.9119
7.9705 | -102.8622
7.9742 | -102.8069
7.9779 | -102.7594
7.9816 | -102.7281
7.9852 | -102.6905
7.9889 | -102.6599
7.9926 | -102.6357
7.9963 | -102.6171
8 | -102.6065
8.0037 | -102.6006
8.0074 | -102.6024
8.0111 | -102.608
8.0148 | -102.6171
8.0184 | -102.6326
8.0221 | -102.6519
8.0258 | -102.675
8.0295 | -102.6971
8.0332 | -102.7162
8.0369 | -102.7312
8.0406 | -102.7457
8.0443 | -102.7609
8.0479 | -102.7736
8.0516 | -102.78
8.0553 | -102.7855
8.059 | -102.793
8.0627 | -102.8019
8.0664 | -102.8115
8.0701 | -102.827
8.0738 | -102.8468
8.0775 | -102.8687
8.0811 | -102.8947
8.0848 | -102.9211
8.0885 | -102.9544
8.0922 | -102.9891
8.0959 | -103.0164
8.0996 | -103.0457
8.1033 | -103.0787
8.107 | -103.1122
8.1106 | -103.1439
8.1143 | -103.1739
8.118 | -103.204
8.1217 | -103.2327
8.1254 | -103.2619
8.1291 | -103.2861
8.1328 | -103.3085
8.1365 | -103.3282
8.1402 | -103.338
8.1438 | -103.3458
8.1475 | -103.3472
8.1512 | -103.3426
8.1549 | -103.3346
8.1586 | -103.3207
8.1623 | -103.2911
8.166 | -103.2504
8.1697 | -103.2094
8.1734 | -103.1701
8.177 | -103.1306
8.1807 | -103.0875
8.1844 | -103.0413
8.1881 | -102.9984
8.1918 | -102.9609
8.1955 | -102.9281
8.1992 | -102.9001
8.2029 | -102.8731
8.2065 | -102.8537
8.2102 | -102.8419
8.2139 | -102.8361
8.2176 | -102.8339
8.2213 | -102.836
8.225 | -102.8416
8.2287 | -102.8481
8.2324 | -102.8568
8.2361 | -102.8704
8.2397 | -102.8823
8.2434 | -102.8979
8.2471 | -102.9227
8.2508 | -102.9521
8.2545 | -102.9826
8.2582 | -103.0166
8.2619 | -103.0577
8.2656 | -103.1009
8.2692 | -103.1477
8.2729 | -103.1913
8.2766 | -103.2313
8.2803 | -103.2712
8.284 | -103.3071
8.2877 | -103.3345
8.2914 | -103.351
8.2951 | -103.3622
8.2988 | -103.3613
8.3024 | -103.3515
8.3061 | -103.3368
8.3098 | -103.3145
8.3135 | -103.2878
8.3172 | -103.2581
8.3209 | -103.2244
8.3246 | -103.1905
8.3283 | -103.1592
8.332 | -103.1314
8.3356 | -103.1099
8.3393 | -103.0918
8.343 | -103.0761
8.3467 | -103.0664
8.3504 | -103.0583
8.3541 | -103.0573
8.3578 | -103.0571
8.3615 | -103.0639
8.3651 | -103.0705
8.3688 | -103.0703
8.3725 | -103.0682
8.3762 | -103.0653
8.3799 | -103.062
8.3836 | -103.0558
8.3873 | -103.0493
8.391 | -103.0427
8.3947 | -103.0372
8.3983 | -103.0312
8.402 | -103.0237
8.4057 | -103.0189
8.4094 | -103.0193
8.4131 | -103.0176
8.4168 | -103.0194
8.4205 | -103.0293
8.4242 | -103.0384
8.4278 | -103.0525
8.4315 | -103.0691
8.4352 | -103.0872
8.4389 | -103.1081
8.4426 | -103.1318
8.4463 | -103.1558
8.45 | -103.1789
8.4537 | -103.1989
8.4574 | -103.215
8.461 | -103.2383
8.4647 | -103.2564
8.4684 | -103.2697
8.4721 | -103.2863
8.4758 | -103.2949
8.4795 | -103.2963
8.4832 | -103.2914
8.4869 | -103.2846
8.4905 | -103.2759
8.4942 | -103.2637
8.4979 | -103.2454
8.5016 | -103.225
8.5053 | -103.2024
8.509 | -103.1767
8.5127 | -103.1493
8.5164 | -103.1245
8.5201 | -103.1068
8.5237 | -103.0886
8.5274 | -103.0739
8.5311 | -103.0651
8.5348 | -103.0567
8.5385 | -103.0539
8.5422 | -103.0552
8.5459 | -103.0576
8.5496 | -103.0657
8.5533 | -103.0747
8.5569 | -103.0874
8.5606 | -103.1031
8.5643 | -103.1196
8.568 | -103.1402
8.5717 | -103.1585
8.5754 | -103.1786
8.5791 | -103.2053
8.5828 | -103.2347
8.5864 | -103.2619
8.5901 | -103.2902
8.5938 | -103.318
8.5975 | -103.3444
8.6012 | -103.3701
8.6049 | -103.3912
8.6086 | -103.4103
8.6123 | -103.424
8.616 | -103.4315
8.6196 | -103.4352
8.6233 | -103.436
8.627 | -103.4332
8.6307 | -103.4251
8.6344 | -103.4112
8.6381 | -103.3962
8.6418 | -103.3812
8.6455 | -103.3601
8.6491 | -103.3336
8.6528 | -103.3078
8.6565 | -103.2803
8.6602 | -103.2547
8.6639 | -103.2362
8.6676 | -103.22
8.6713 | -103.2088
8.675 | -103.2059
8.6787 | -103.209
8.6823 | -103.2178
8.686 | -103.2327
8.6897 | -103.2536
8.6934 | -103.2808
8.6971 | -103.3123
8.7008 | -103.348
8.7045 | -103.3816
8.7082 | -103.4135
8.7118 | -103.4493
8.7155 | -103.4868
8.7192 | -103.5232
8.7229 | -103.5571
8.7266 | -103.5889
8.7303 | -103.6143
8.734 | -103.643
8.7377 | -103.6629
8.7414 | -103.6773
8.745 | -103.6882
8.7487 | -103.6915
8.7524 | -103.6925
8.7561 | -103.69
8.7598 | -103.6868
8.7635 | -103.6834
8.7672 | -103.6779
8.7709 | -103.6651
8.7745 | -103.6482
8.7782 | -103.6259
8.7819 | -103.605
8.7856 | -103.5879
8.7893 | -103.5675
8.793 | -103.545
8.7967 | -103.5257
8.8004 | -103.5054
8.8041 | -103.4846
8.8077 | -103.4648
8.8114 | -103.442
8.8151 | -103.425
8.8188 | -103.4117
8.8225 | -103.4013
8.8262 | -103.3917
8.8299 | -103.3837
8.8336 | -103.377
8.8373 | -103.3708
8.8409 | -103.3607
8.8446 | -103.35
8.8483 | -103.3456
8.852 | -103.3425
8.8557 | -103.3384
8.8594 | -103.3369
8.8631 | -103.3355
8.8668 | -103.3346
8.8704 | -103.3351
8.8741 | -103.3376
8.8778 | -103.3444
8.8815 | -103.3496
8.8852 | -103.3487
8.8889 | -103.3503
8.8926 | -103.3524
8.8963 | -103.3519
8.9 | -103.3524
8.9036 | -103.355
8.9073 | -103.3549
8.911 | -103.3483
8.9147 | -103.3377
8.9184 | -103.3275
8.9221 | -103.3141
8.9258 | -103.3006
8.9295 | -103.2911
8.9331 | -103.2847
8.9368 | -103.281
8.9405 | -103.2815
8.9442 | -103.2782
8.9479 | -103.2728
8.9516 | -103.2697
8.9553 | -103.2709
8.959 | -103.2756
8.9627 | -103.2829
8.9663 | -103.2941
8.97 | -103.309
8.9737 | -103.324
8.9774 | -103.3354
8.9811 | -103.3492
8.9848 | -103.3635
8.9885 | -103.3792
8.9922 | -103.3989
8.9959 | -103.4212
8.9995 | -103.4418
9.0032 | -103.4584
9.0069 | -103.469
9.0106 | -103.4729
9.0143 | -103.4784
9.018 | -103.4783
9.0217 | -103.4737
9.0254 | -103.4616
9.029 | -103.4466
9.0327 | -103.4254
9.0364 | -103.4007
9.0401 | -103.3768
9.0438 | -103.3594
9.0475 | -103.3422
9.0512 | -103.3256
9.0549 | -103.3176
9.0586 | -103.3107
9.0622 | -103.299
9.0659 | -103.2898
9.0696 | -103.2865
9.0733 | -103.2846
9.077 | -103.284
9.0807 | -103.2854
9.0844 | -103.2872
9.0881 | -103.289
9.0917 | -103.2906
9.0954 | -103.2873
9.0991 | -103.2802
9.1028 | -103.2753
9.1065 | -103.2731
9.1102 | -103.2711
9.1139 | -103.267
9.1176 | -103.2649
9.1213 | -103.2595
9.1249 | -103.2516
9.1286 | -103.2429
9.1323 | -103.2379
9.136 | -103.2377
9.1397 | -103.2432
9.1434 | -103.253
9.1471 | -103.2615
9.1508 | -103.2695
9.1544 | -103.2777
9.1581 | -103.2909
9.1618 | -103.305
9.1655 | -103.3181
9.1692 | -103.3358
9.1729 | -103.3524
9.1766 | -103.3713
9.1803 | -103.3856
9.184 | -103.4035
9.1876 | -103.4253
9.1913 | -103.4415
9.195 | -103.4541
9.1987 | -103.4687
9.2024 | -103.4808
9.2061 | -103.4893
9.2098 | -103.496
9.2135 | -103.5022
9.2172 | -103.5046
9.2208 | -103.5004
9.2245 | -103.4949
9.2282 | -103.4887
9.2319 | -103.4819
9.2356 | -103.4757
9.2393 | -103.4689
9.243 | -103.461
9.2467 | -103.4528
9.2503 | -103.4484
9.254 | -103.446
9.2577 | -103.4509
9.2614 | -103.4575
9.2651 | -103.464
9.2688 | -103.4712
9.2725 | -103.4785
9.2762 | -103.4837
9.2799 | -103.4919
9.2835 | -103.5019
9.2872 | -103.5149
9.2909 | -103.5334
9.2946 | -103.5491
9.2983 | -103.5614
9.302 | -103.5759
9.3057 | -103.5914
9.3094 | -103.605
9.313 | -103.6175
9.3167 | -103.6378
9.3204 | -103.6581
9.3241 | -103.6773
9.3278 | -103.6986
9.3315 | -103.7225
9.3352 | -103.7438
9.3389 | -103.7641
9.3426 | -103.7859
9.3462 | -103.8119
9.3499 | -103.8342
9.3536 | -103.8551
9.3573 | -103.8782
9.361 | -103.8987
9.3647 | -103.9145
9.3684 | -103.9316
9.3721 | -103.945
9.3757 | -103.952
9.3794 | -103.9564
9.3831 | -103.9591
9.3868 | -103.9606
9.3905 | -103.954
9.3942 | -103.9436
9.3979 | -103.9319
9.4016 | -103.9234
9.4053 | -103.9079
9.4089 | -103.8934
9.4126 | -103.8841
9.4163 | -103.874
9.42 | -103.8617
9.4237 | -103.8493
9.4274 | -103.8356
9.4311 | -103.8298
9.4348 | -103.8301
9.4385 | -103.8278
9.4421 | -103.823
9.4458 | -103.8268
9.4495 | -103.8354
9.4532 | -103.8424
9.4569 | -103.8503
9.4606 | -103.8659
9.4643 | -103.8869
9.468 | -103.9106
9.4716 | -103.9359
9.4753 | -103.956
9.479 | -103.9725
9.4827 | -103.9873
9.4864 | -104.0047
9.4901 | -104.0242
9.4938 | -104.0408
9.4975 | -104.0537
9.5012 | -104.0661
9.5048 | -104.0765
9.5085 | -104.0843
9.5122 | -104.0929
9.5159 | -104.1024
9.5196 | -104.1109
9.5233 | -104.1158
9.527 | -104.128
9.5307 | -104.1451
9.5343 | -104.1643
9.538 | -104.1839
9.5417 | -104.2046
9.5454 | -104.2223
9.5491 | -104.2324
9.5528 | -104.242
9.5565 | -104.2502
9.5602 | -104.2569
9.5639 | -104.262
9.5675 | -104.2636
9.5712 | -104.2622
9.5749 | -104.2549
9.5786 | -104.2552
9.5823 | -104.2538
9.586 | -104.2456
9.5897 | -104.2331
9.5934 | -104.224
9.597 | -104.2179
9.6007 | -104.2092
9.6044 | -104.1989
9.6081 | -104.1856
9.6118 | -104.1719
9.6155 | -104.1615
9.6192 | -104.1495
9.6229 | -104.1322
9.6266 | -104.1171
9.6302 | -104.1024
9.6339 | -104.088
9.6376 | -104.076
9.6413 | -104.0658
9.645 | -104.0564
9.6487 | -104.0479
9.6524 | -104.0367
9.6561 | -104.0262
9.6598 | -104.0164
9.6634 | -104.0066
9.6671 | -103.994
9.6708 | -103.9801
9.6745 | -103.9649
9.6782 | -103.9486
9.6819 | -103.9372
9.6856 | -103.9291
9.6893 | -103.9227
9.6929 | -103.9166
9.6966 | -103.9125
9.7003 | -103.9061
9.704 | -103.9011
9.7077 | -103.9007
9.7114 | -103.9039
9.7151 | -103.907
9.7188 | -103.913
9.7225 | -103.9211
9.7261 | -103.9318
9.7298 | -103.9454
9.7335 | -103.9632
9.7372 | -103.98
9.7409 | -103.9969
9.7446 | -104.013
9.7483 | -104.0316
9.752 | -104.0527
9.7556 | -104.0737
9.7593 | -104.0938
9.763 | -104.1126
9.7667 | -104.132
9.7704 | -104.1501
9.7741 | -104.1642
9.7778 | -104.1769
9.7815 | -104.1856
9.7852 | -104.1932
9.7888 | -104.1996
9.7925 | -104.2085
9.7962 | -104.2182
9.7999 | -104.2248
9.8036 | -104.2292
9.8073 | -104.2281
9.811 | -104.2266
9.8147 | -104.2272
9.8183 | -104.2295
9.822 | -104.2355
9.8257 | -104.2373
9.8294 | -104.2368
9.8331 | -104.2357
9.8368 | -104.235
9.8405 | -104.2333
9.8442 | -104.2325
9.8479 | -104.2339
9.8515 | -104.2318
9.8552 | -104.2261
9.8589 | -104.225
9.8626 | -104.2256
9.8663 | -104.2194
9.87 | -104.2137
9.8737 | -104.2091
9.8774 | -104.2048
9.8811 | -104.2041
9.8847 | -104.2011
9.8884 | -104.1978
9.8921 | -104.198
9.8958 | -104.199
9.8995 | -104.1993
9.9032 | -104.1992
9.9069 | -104.2005
9.9106 | -104.2005
9.9142 | -104.1925
9.9179 | -104.1871
9.9216 | -104.1838
9.9253 | -104.1835
9.929 | -104.1817
9.9327 | -104.1784
9.9364 | -104.1738
9.9401 | -104.1722
9.9438 | -104.1747
9.9474 | -104.1813
9.9511 | -104.1909
9.9548 | -104.1988
9.9585 | -104.2064
9.9622 | -104.2169
9.9659 | -104.23
9.9696 | -104.2465
9.9733 | -104.2652
9.9769 | -104.2869
9.9806 | -104.3096
9.9843 | -104.3329
9.988 | -104.3547
9.9917 | -104.3768
9.9954 | -104.401
9.9991 | -104.4254
10.0028 | -104.4445
10.0065 | -104.4621
10.0101 | -104.4843
10.0138 | -104.5009
10.0175 | -104.5089
10.0212 | -104.5138
10.0249 | -104.5141
10.0286 | -104.5109
10.0323 | -104.5029
10.036 | -104.4887
10.0396 | -104.469
10.0433 | -104.4494
10.047 | -104.4295
10.0507 | -104.4046
10.0544 | -104.3821
10.0581 | -104.3684
10.0618 | -104.3502
10.0655 | -104.3329
10.0692 | -104.3178
10.0728 | -104.3078
10.0765 | -104.3046
10.0802 | -104.3021
10.0839 | -104.2973
10.0876 | -104.2957
10.0913 | -104.2916
10.095 | -104.2891
10.0987 | -104.2884
10.1024 | -104.2869
10.106 | -104.2791
10.1097 | -104.2551
10.1134 | -104.2375
10.1171 | -104.217
10.1208 | -104.1926
10.1245 | -104.165
10.1282 | -104.1297
10.1319 | -104.0933
10.1355 | -104.0562
10.1392 | -104.0163
10.1429 | -103.975
10.1466 | -103.937
10.1503 | -103.8992
10.154 | -103.8586
10.1577 | -103.8195
10.1614 | -103.7866
10.1651 | -103.7603
10.1687 | -103.7378
10.1724 | -103.7222
10.1761 | -103.7148
10.1798 | -103.7135
10.1835 | -103.7173
10.1872 | -103.7287
10.1909 | -103.7433
10.1946 | -103.7606
10.1982 | -103.7871
10.2019 | -103.8128
10.2056 | -103.8376
10.2093 | -103.868
10.213 | -103.8998
10.2167 | -103.9317
10.2204 | -103.9602
10.2241 | -103.9859
10.2278 | -104.0143
10.2314 | -104.0415
10.2351 | -104.0606
10.2388 | -104.0755
10.2425 | -104.086
10.2462 | -104.0934
10.2499 | -104.0931
10.2536 | -104.0931
10.2573 | -104.091
10.2609 | -104.0826
10.2646 | -104.0762
10.2683 | -104.0699
10.272 | -104.0618
10.2757 | -104.0534
10.2794 | -104.0457
10.2831 | -104.0366
10.2868 | -104.0267
10.2905 | -104.0131
10.2941 | -104.0032
10.2978 | -103.996
10.3015 | -103.9957
10.3052 | -104.0014
10.3089 | -104.0068
10.3126 | -104.013
10.3163 | -104.0236
10.32 | -104.0356
10.3237 | -104.0517
10.3273 | -104.073
10.331 | -104.0985
10.3347 | -104.1237
10.3384 | -104.1494
10.3421 | -104.1787
10.3458 | -104.2111
10.3495 | -104.2412
10.3532 | -104.2708
10.3568 | -104.2969
10.3605 | -104.3188
10.3642 | -104.3366
10.3679 | -104.3493
10.3716 | -104.3589
10.3753 | -104.3634
10.379 | -104.3629
10.3827 | -104.3621
10.3864 | -104.3605
10.39 | -104.3596
10.3937 | -104.351
10.3974 | -104.3383
10.4011 | -104.3294
10.4048 | -104.3194
10.4085 | -104.3073
10.4122 | -104.2963
10.4159 | -104.2871
10.4195 | -104.2785
10.4232 | -104.2703
10.4269 | -104.2655
10.4306 | -104.2652
10.4343 | -104.2677
10.438 | -104.2724
10.4417 | -104.2829
10.4454 | -104.2952
10.4491 | -104.3073
10.4527 | -104.3157
10.4564 | -104.3255
10.4601 | -104.3399
10.4638 | -104.3515
10.4675 | -104.3612
10.4712 | -104.3729
10.4749 | -104.3801
10.4786 | -104.3839
10.4823 | -104.389
10.4859 | -104.393
10.4896 | -104.3942
10.4933 | -104.3919
10.497 | -104.3844
10.5007 | -104.3752
10.5044 | -104.3696
10.5081 | -104.3612
10.5118 | -104.3529
10.5154 | -104.3463
10.5191 | -104.3435
10.5228 | -104.3434
10.5265 | -104.3428
10.5302 | -104.3432
10.5339 | -104.3466
10.5376 | -104.349
10.5413 | -104.3534
10.545 | -104.363
10.5486 | -104.3733
10.5523 | -104.3835
10.556 | -104.3945
10.5597 | -104.4058
10.5634 | -104.4203
10.5671 | -104.4331
10.5708 | -104.442
10.5745 | -104.4501
10.5781 | -104.4562
10.5818 | -104.4607
10.5855 | -104.4596
10.5892 | -104.4539
10.5929 | -104.4488
10.5966 | -104.4418
10.6003 | -104.4347
10.604 | -104.4307
10.6077 | -104.4261
10.6113 | -104.4211
10.615 | -104.4204
10.6187 | -104.4221
10.6224 | -104.4235
10.6261 | -104.4306
10.6298 | -104.4414
10.6335 | -104.4481
10.6372 | -104.4582
10.6408 | -104.4884
10.6445 | -104.5093
10.6482 | -104.5314
10.6519 | -104.5568
10.6556 | -104.5828
10.6593 | -104.6029
10.663 | -104.6264
10.6667 | -104.6487
10.6704 | -104.6646
10.674 | -104.6784
10.6777 | -104.692
10.6814 | -104.7004
10.6851 | -104.7061
10.6888 | -104.7092
10.6925 | -104.7105
10.6962 | -104.7049
10.6999 | -104.6987
10.7035 | -104.6893
10.7072 | -104.6757
10.7109 | -104.6655
10.7146 | -104.6579
10.7183 | -104.6453
10.722 | -104.6321
10.7257 | -104.6214
10.7294 | -104.6096
10.7331 | -104.5962
10.7367 | -104.5805
10.7404 | -104.5666
10.7441 | -104.5574
10.7478 | -104.5456
10.7515 | -104.5326
10.7552 | -104.5195
10.7589 | -104.504
10.7626 | -104.4905
10.7663 | -104.4767
10.7699 | -104.4623
10.7736 | -104.4473
10.7773 | -104.4338
10.781 | -104.4214
10.7847 | -104.409
10.7884 | -104.3964
10.7921 | -104.3828
10.7958 | -104.3694
10.7994 | -104.3581
10.8031 | -104.3521
10.8068 | -104.3472
10.8105 | -104.3443
10.8142 | -104.3433
10.8179 | -104.349
10.8216 | -104.3573
10.8253 | -104.3687
10.829 | -104.3823
10.8326 | -104.398
10.8363 | -104.415
10.84 | -104.4403
10.8437 | -104.4661
10.8474 | -104.4904
10.8511 | -104.5206
10.8548 | -104.5522
10.8585 | -104.5805
10.8621 | -104.6083
10.8658 | -104.6345
10.8695 | -104.6578
10.8732 | -104.6785
10.8769 | -104.6996
10.8806 | -104.7193
10.8843 | -104.739
10.888 | -104.7553
10.8917 | -104.7692
10.8953 | -104.783
10.899 | -104.7916
10.9027 | -104.7982
10.9064 | -104.8038
10.9101 | -104.8103
10.9138 | -104.8208
10.9175 | -104.8336
10.9212 | -104.8403
10.9248 | -104.8486
10.9285 | -104.8579
10.9322 | -104.8657
10.9359 | -104.8741
10.9396 | -104.8806
10.9433 | -104.8892
10.947 | -104.9029
10.9507 | -104.9143
10.9544 | -104.9166
10.958 | -104.9139
10.9617 | -104.9152
10.9654 | -104.9197
10.9691 | -104.9223
10.9728 | -104.9208
10.9765 | -104.9153
10.9802 | -104.906
10.9839 | -104.8962
10.9876 | -104.8864
10.9912 | -104.8755
10.9949 | -104.8633
10.9986 | -104.8451
11.0023 | -104.8231
11.006 | -104.8029
11.0097 | -104.7777
11.0134 | -104.754
11.0171 | -104.7296
11.0207 | -104.7047
11.0244 | -104.6796
11.0281 | -104.6518
11.0318 | -104.6219
11.0355 | -104.5946
11.0392 | -104.5686
11.0429 | -104.5449
11.0466 | -104.524
11.0503 | -104.4991
11.0539 | -104.4765
11.0576 | -104.4621
11.0613 | -104.4518
11.065 | -104.4448
11.0687 | -104.4422
11.0724 | -104.4434
11.0761 | -104.4505
11.0798 | -104.4641
11.0834 | -104.4795
11.0871 | -104.5009
11.0908 | -104.5277
11.0945 | -104.5566
11.0982 | -104.585
11.1019 | -104.6141
11.1056 | -104.6416
11.1093 | -104.6697
11.113 | -104.7017
11.1166 | -104.7297
11.1203 | -104.749
11.124 | -104.7676
11.1277 | -104.7854
11.1314 | -104.8041
11.1351 | -104.8212
11.1388 | -104.8344
11.1425 | -104.8423
11.1462 | -104.8471
11.1498 | -104.8471
11.1535 | -104.8419
11.1572 | -104.8367
11.1609 | -104.8323
11.1646 | -104.825
11.1683 | -104.8169
11.172 | -104.7973
11.1757 | -104.7861
11.1793 | -104.7725
11.183 | -104.7561
11.1867 | -104.7398
11.1904 | -104.7257
11.1941 | -104.7143
11.1978 | -104.7003
11.2015 | -104.687
11.2052 | -104.6753
11.2089 | -104.6665
11.2125 | -104.6648
11.2162 | -104.6637
11.2199 | -104.6605
11.2236 | -104.6579
11.2273 | -104.6573
11.231 | -104.661
11.2347 | -104.667
11.2384 | -104.6747
11.242 | -104.6884
11.2457 | -104.7055
11.2494 | -104.7216
11.2531 | -104.7373
11.2568 | -104.7518
11.2605 | -104.7671
11.2642 | -104.7813
11.2679 | -104.8003
11.2716 | -104.8178
11.2752 | -104.8328
11.2789 | -104.8461
11.2826 | -104.8572
11.2863 | -104.8675
11.29 | -104.8787
11.2937 | -104.885
11.2974 | -104.8898
11.3011 | -104.8925
11.3047 | -104.8928
11.3084 | -104.8955
11.3121 | -104.8957
11.3158 | -104.894
11.3195 | -104.8917
11.3232 | -104.8841
11.3269 | -104.8728
11.3306 | -104.8594
11.3343 | -104.8444
11.3379 | -104.831
11.3416 | -104.8118
11.3453 | -104.7933
11.349 | -104.7728
11.3527 | -104.755
11.3564 | -104.7337
11.3601 | -104.7109
11.3638 | -104.6933
11.3674 | -104.6758
11.3711 | -104.6563
11.3748 | -104.6409
11.3785 | -104.6259
11.3822 | -104.6126
11.3859 | -104.6051
11.3896 | -104.5972
11.3933 | -104.595
11.397 | -104.5944
11.4006 | -104.5975
11.4043 | -104.6027
11.408 | -104.6053
11.4117 | -104.6136
11.4154 | -104.627
11.4191 | -104.6415
11.4228 | -104.6603
11.4265 | -104.6841
11.4302 | -104.7091
11.4338 | -104.7294
11.4375 | -104.7526
11.4412 | -104.7741
11.4449 | -104.7988
11.4486 | -104.8257
11.4523 | -104.848
11.456 | -104.8674
11.4597 | -104.8856
11.4633 | -104.9002
11.467 | -104.9113
11.4707 | -104.9214
11.4744 | -104.9278
11.4781 | -104.9313
11.4818 | -104.933
11.4855 | -104.9305
11.4892 | -104.9274
11.4929 | -104.924
11.4965 | -104.9183
11.5002 | -104.9102
11.5039 | -104.8971
11.5076 | -104.8836
11.5113 | -104.8663
11.515 | -104.8505
11.5187 | -104.8325
11.5224 | -104.8168
11.526 | -104.8023
11.5297 | -104.7836
11.5334 | -104.7666
11.5371 | -104.7521
11.5408 | -104.7356
11.5445 | -104.7167
11.5482 | -104.7014
11.5519 | -104.6852
11.5556 | -104.6754
11.5592 | -104.6683
11.5629 | -104.6611
11.5666 | -104.6538
11.5703 | -104.6495
11.574 | -104.6483
11.5777 | -104.6458
11.5814 | -104.6438
11.5851 | -104.6475
11.5888 | -104.6501
11.5924 | -104.6495
11.5961 | -104.6486
11.5998 | -104.6531
11.6035 | -104.6602
11.6072 | -104.6686
11.6109 | -104.6742
11.6146 | -104.683
11.6183 | -104.6947
11.6219 | -104.7065
11.6256 | -104.7234
11.6293 | -104.7433
11.633 | -104.7604
11.6367 | -104.7748
11.6404 | -104.7924
11.6441 | -104.8115
11.6478 | -104.8309
11.6515 | -104.8497
11.6551 | -104.8652
11.6588 | -104.8797
11.6625 | -104.8934
11.6662 | -104.9085
11.6699 | -104.9209
11.6736 | -104.9275
11.6773 | -104.9362
11.681 | -104.9425
11.6846 | -104.9433
11.6883 | -104.9419
11.692 | -104.9376
11.6957 | -104.9363
11.6994 | -104.9317
11.7031 | -104.9227
11.7068 | -104.9168
11.7105 | -104.9116
11.7142 | -104.9075
11.7178 | -104.8961
11.7215 | -104.886
11.7252 | -104.8728
11.7289 | -104.8569
11.7326 | -104.8412
11.7363 | -104.8278
11.74 | -104.8129
11.7437 | -104.7959
11.7473 | -104.7749
11.751 | -104.7555
11.7547 | -104.7389
11.7584 | -104.7241
11.7621 | -104.7079
11.7658 | -104.6916
11.7695 | -104.6717
11.7732 | -104.6514
11.7769 | -104.6318
11.7805 | -104.6126
11.7842 | -104.5935
11.7879 | -104.5772
11.7916 | -104.564
11.7953 | -104.548
11.799 | -104.5296
11.8027 | -104.5159
11.8064 | -104.5019
11.8101 | -104.487
11.8137 | -104.4731
11.8174 | -104.4621
11.8211 | -104.4534
11.8248 | -104.4467
11.8285 | -104.4414
11.8322 | -104.4398
11.8359 | -104.4393
11.8396 | -104.4418
11.8432 | -104.451
11.8469 | -104.4606
11.8506 | -104.4728
11.8543 | -104.4869
11.858 | -104.504
11.8617 | -104.526
11.8654 | -104.5474
11.8691 | -104.5677
11.8728 | -104.586
11.8764 | -104.607
11.8801 | -104.629
11.8838 | -104.6493
11.8875 | -104.6669
11.8912 | -104.6849
11.8949 | -104.7012
11.8986 | -104.7167
11.9023 | -104.7296
11.9059 | -104.7427
11.9096 | -104.7486
11.9133 | -104.7543
11.917 | -104.7604
11.9207 | -104.7637
11.9244 | -104.7623
11.9281 | -104.7564
11.9318 | -104.7503
11.9355 | -104.746
11.9391 | -104.7429
11.9428 | -104.7389
11.9465 | -104.7351
11.9502 | -104.7314
11.9539 | -104.7283
11.9576 | -104.7242
11.9613 | -104.7155
11.965 | -104.7056
11.9686 | -104.7005
11.9723 | -104.7027
11.976 | -104.7051
11.9797 | -104.7097
11.9834 | -104.7122
11.9871 | -104.7173
11.9908 | -104.7228
11.9945 | -104.7288
11.9982 | -104.7373
12.0018 | -104.7463
12.0055 | -104.7577
12.0092 | -104.7696
12.0129 | -104.7837
12.0166 | -104.7977
12.0203 | -104.8118
12.024 | -104.8248
12.0277 | -104.838
12.0314 | -104.8532
12.035 | -104.8685
12.0387 | -104.8826
12.0424 | -104.8958
12.0461 | -104.9073
12.0498 | -104.9179
12.0535 | -104.9248
12.0572 | -104.9306
12.0609 | -104.9387
12.0645 | -104.9446
12.0682 | -104.9446
12.0719 | -104.9434
12.0756 | -104.9416
12.0793 | -104.941
12.083 | -104.9433
12.0867 | -104.9396
12.0904 | -104.9386
12.0941 | -104.9412
12.0977 | -104.9421
12.1014 | -104.9384
12.1051 | -104.9337
12.1088 | -104.929
12.1125 | -104.9238
12.1162 | -104.9169
12.1199 | -104.9062
12.1236 | -104.897
12.1272 | -104.8863
12.1309 | -104.8699
12.1346 | -104.8596
12.1383 | -104.8513
12.142 | -104.8395
12.1457 | -104.8245
12.1494 | -104.8135
12.1531 | -104.8053
12.1568 | -104.8012
12.1604 | -104.7975
12.1641 | -104.7968
12.1678 | -104.7985
12.1715 | -104.8038
12.1752 | -104.8101
12.1789 | -104.8195
12.1826 | -104.8312
12.1863 | -104.8444
12.1899 | -104.8582
12.1936 | -104.8774
12.1973 | -104.8999
12.201 | -104.9246
12.2047 | -104.945
12.2084 | -104.9647
12.2121 | -104.9852
12.2158 | -105.0085
12.2195 | -105.0322
12.2231 | -105.0538
12.2268 | -105.0733
12.2305 | -105.0913
12.2342 | -105.106
12.2379 | -105.104
12.2416 | -105.1153
12.2453 | -105.1286
12.249 | -105.1386
12.2527 | -105.1452
12.2563 | -105.1479
12.26 | -105.1481
12.2637 | -105.1508
12.2674 | -105.1482
12.2711 | -105.1411
12.2748 | -105.1361
12.2785 | -105.1352
12.2822 | -105.1329
12.2858 | -105.129
12.2895 | -105.1209
12.2932 | -105.1149
12.2969 | -105.1106
12.3006 | -105.1028
12.3043 | -105.0952
12.308 | -105.0899
12.3117 | -105.0797
12.3154 | -105.0677
12.319 | -105.061
12.3227 | -105.0543
12.3264 | -105.0481
12.3301 | -105.0416
12.3338 | -105.0357
12.3375 | -105.0314
12.3412 | -105.0253
12.3449 | -105.0231
12.3485 | -105.0213
12.3522 | -105.0205
12.3559 | -105.014
12.3596 | -105.0103
12.3633 | -105.0081
12.367 | -105.0078
12.3707 | -105.0082
12.3744 | -105.0082
12.3781 | -105.0075
12.3817 | -105.0105
12.3854 | -105.0133
12.3891 | -105.0168
12.3928 | -105.024
12.3965 | -105.0327
12.4002 | -105.0424
12.4039 | -105.0509
12.4076 | -105.059
12.4112 | -105.0703
12.4149 | -105.083
12.4186 | -105.0977
12.4223 | -105.1104
12.426 | -105.1212
12.4297 | -105.1339
12.4334 | -105.1479
12.4371 | -105.157
12.4408 | -105.1664
12.4444 | -105.1728
12.4481 | -105.1768
12.4518 | -105.1804
12.4555 | -105.1804
12.4592 | -105.1792
12.4629 | -105.178
12.4666 | -105.177
12.4703 | -105.1767
12.474 | -105.1726
12.4776 | -105.167
12.4813 | -105.1654
12.485 | -105.1643
12.4887 | -105.1633
12.4924 | -105.1659
12.4961 | -105.168
12.4998 | -105.1664
12.5035 | -105.1693
12.5071 | -105.1713
12.5108 | -105.1756
12.5145 | -105.181
12.5182 | -105.1891
12.5219 | -105.1997
12.5256 | -105.2084
12.5293 | -105.2165
12.533 | -105.224
12.5367 | -105.233
12.5403 | -105.2401
12.544 | -105.2452
12.5477 | -105.2491
12.5514 | -105.2522
12.5551 | -105.2575
12.5588 | -105.2647
12.5625 | -105.2715
12.5662 | -105.2769
12.5698 | -105.2819
12.5735 | -105.2848
12.5772 | -105.2877
12.5809 | -105.2901
12.5846 | -105.2968
12.5883 | -105.302
12.592 | -105.3004
12.5957 | -105.3009
12.5994 | -105.3002
12.603 | -105.2988
12.6067 | -105.3004
12.6104 | -105.2997
12.6141 | -105.2969
12.6178 | -105.3
12.6215 | -105.3003
12.6252 | -105.2972
12.6289 | -105.2932
12.6326 | -105.2878
12.6362 | -105.2825
12.6399 | -105.2757
12.6436 | -105.2694
12.6473 | -105.2602
12.651 | -105.2502
12.6547 | -105.2454
12.6584 | -105.2433
12.6621 | -105.2393
12.6657 | -105.2341
12.6694 | -105.2291
12.6731 | -105.2183
12.6768 | -105.21
12.6805 | -105.205
12.6842 | -105.1972
12.6879 | -105.1896
12.6916 | -105.1816
12.6953 | -105.177
12.6989 | -105.1716
12.7026 | -105.166
12.7063 | -105.1576
12.71 | -105.1491
12.7137 | -105.1429
12.7174 | -105.137
12.7211 | -105.1325
12.7248 | -105.1295
12.7284 | -105.1253
12.7321 | -105.1193
12.7358 | -105.1118
12.7395 | -105.1044
12.7432 | -105.0962
12.7469 | -105.0905
12.7506 | -105.0863
12.7543 | -105.0817
12.758 | -105.0761
12.7616 | -105.0713
12.7653 | -105.067
12.769 | -105.0643
12.7727 | -105.0623
12.7764 | -105.061
12.7801 | -105.0556
12.7838 | -105.0503
12.7875 | -105.0451
12.7911 | -105.036
12.7948 | -105.0316
12.7985 | -105.0324
12.8022 | -105.0308
12.8059 | -105.0279
12.8096 | -105.0312
12.8133 | -105.0347
12.817 | -105.0384
12.8207 | -105.0426
12.8243 | -105.0453
12.828 | -105.046
12.8317 | -105.0494
12.8354 | -105.0554
12.8391 | -105.0577
12.8428 | -105.0561
12.8465 | -105.0529
12.8502 | -105.0542
12.8538 | -105.0539
12.8575 | -105.0555
12.8612 | -105.0569
12.8649 | -105.0553
12.8686 | -105.052
12.8723 | -105.0539
12.876 | -105.0558
12.8797 | -105.055
12.8834 | -105.0539
12.887 | -105.054
12.8907 | -105.059
12.8944 | -105.0634
12.8981 | -105.0665
12.9018 | -105.0689
12.9055 | -105.0722
12.9092 | -105.0758
12.9129 | -105.0824
12.9166 | -105.0921
12.9202 | -105.1043
12.9239 | -105.1158
12.9276 | -105.1282
12.9313 | -105.1414
12.935 | -105.1549
12.9387 | -105.1688
12.9424 | -105.1837
12.9461 | -105.2021
12.9497 | -105.2204
12.9534 | -105.2383
12.9571 | -105.2562
12.9608 | -105.2734
12.9645 | -105.294
12.9682 | -105.3164
12.9719 | -105.3371
12.9756 | -105.3562
12.9793 | -105.3793
12.9829 | -105.3983
12.9866 | -105.4164
12.9903 | -105.4366
12.994 | -105.4546
12.9977 | -105.4683
13.0014 | -105.4847
13.0051 | -105.5018
13.0088 | -105.5174
13.0124 | -105.5311
13.0161 | -105.5458
13.0198 | -105.5582
13.0235 | -105.568
13.0272 | -105.5765
13.0309 | -105.5835
13.0346 | -105.5889
13.0383 | -105.5926
13.042 | -105.5957
13.0456 | -105.5986
13.0493 | -105.5992
13.053 | -105.5982
13.0567 | -105.599
13.0604 | -105.5932
13.0641 | -105.5867
13.0678 | -105.5812
13.0715 | -105.5762
13.0752 | -105.5716
13.0788 | -105.5631
13.0825 | -105.5563
13.0862 | -105.5519
13.0899 | -105.5491
13.0936 | -105.5434
13.0973 | -105.5393
13.101 | -105.5336
13.1047 | -105.5281
13.1083 | -105.523
13.112 | -105.5213
13.1157 | -105.5206
13.1194 | -105.5204
13.1231 | -105.5215
13.1268 | -105.5252
13.1305 | -105.5271
13.1342 | -105.5271
13.1379 | -105.5321
13.1415 | -105.5401
13.1452 | -105.5454
13.1489 | -105.5499
13.1526 | -105.5521
13.1563 | -105.5536
13.16 | -105.5535
13.1637 | -105.554
13.1674 | -105.554
13.171 | -105.5546
13.1747 | -105.5512
13.1784 | -105.5474
13.1821 | -105.5469
13.1858 | -105.5393
13.1895 | -105.5304
13.1932 | -105.5204
13.1969 | -105.5133
13.2006 | -105.508
13.2042 | -105.4993
13.2079 | -105.491
13.2116 | -105.4812
13.2153 | -105.4717
13.219 | -105.4624
13.2227 | -105.4523
13.2264 | -105.4377
13.2301 | -105.4259
13.2337 | -105.4142
13.2374 | -105.4014
13.2411 | -105.3905
13.2448 | -105.3785
13.2485 | -105.3669
13.2522 | -105.3562
13.2559 | -105.3456
13.2596 | -105.335
13.2633 | -105.3248
13.2669 | -105.3177
13.2706 | -105.3098
13.2743 | -105.3011
13.278 | -105.2936
13.2817 | -105.2845
13.2854 | -105.2744
13.2891 | -105.2663
13.2928 | -105.2604
13.2965 | -105.2542
13.3001 | -105.2533
13.3038 | -105.2486
13.3075 | -105.2458
13.3112 | -105.2412
13.3149 | -105.2364
13.3186 | -105.2326
13.3223 | -105.2315
13.326 | -105.2325
13.3296 | -105.23
13.3333 | -105.2286
13.337 | -105.2333
13.3407 | -105.2384
13.3444 | -105.2437
13.3481 | -105.2479
13.3518 | -105.2523
13.3555 | -105.2587
13.3592 | -105.2644
13.3628 | -105.2705
13.3665 | -105.2773
13.3702 | -105.2828
13.3739 | -105.2899
13.3776 | -105.3024
13.3813 | -105.3099
13.385 | -105.3143
13.3887 | -105.3174
13.3923 | -105.3211
13.396 | -105.329
13.3997 | -105.331
13.4034 | -105.3326
13.4071 | -105.3326
13.4108 | -105.3307
13.4145 | -105.3291
13.4182 | -105.3294
13.4219 | -105.3251
13.4255 | -105.319
13.4292 | -105.3137
13.4329 | -105.3082
13.4366 | -105.305
13.4403 | -105.3031
13.444 | -105.3004
13.4477 | -105.2937
13.4514 | -105.2868
13.455 | -105.2834
13.4587 | -105.2819
13.4624 | -105.2778
13.4661 | -105.2764
13.4698 | -105.2771
13.4735 | -105.2804
13.4772 | -105.2849
13.4809 | -105.2887
13.4846 | -105.29
13.4882 | -105.2925
13.4919 | -105.2953
13.4956 | -105.2967
13.4993 | -105.2985
13.503 | -105.2988
13.5067 | -105.3023
13.5104 | -105.3034
13.5141 | -105.3031
13.5177 | -105.3048
13.5214 | -105.3042
13.5251 | -105.3014
13.5288 | -105.2968
13.5325 | -105.29
13.5362 | -105.2849
13.5399 | -105.2803
13.5436 | -105.2768
13.5473 | -105.2709
13.5509 | -105.2611
13.5546 | -105.2504
13.5583 | -105.24
13.562 | -105.2335
13.5657 | -105.2271
13.5694 | -105.2203
13.5731 | -105.2123
13.5768 | -105.2066
13.5805 | -105.2025
13.5841 | -105.1957
13.5878 | -105.1889
13.5915 | -105.1828
13.5952 | -105.1821
13.5989 | -105.1817
13.6026 | -105.1801
13.6063 | -105.1824
13.61 | -105.1876
13.6136 | -105.1909
13.6173 | -105.1977
13.621 | -105.2053
13.6247 | -105.2119
13.6284 | -105.2209
13.6321 | -105.2307
13.6358 | -105.2406
13.6395 | -105.2509
13.6432 | -105.2639
13.6468 | -105.2786
13.6505 | -105.2887
13.6542 | -105.2987
13.6579 | -105.316
13.6616 | -105.3323
13.6653 | -105.3466
13.669 | -105.3601
13.6727 | -105.3733
13.6763 | -105.3862
13.68 | -105.4009
13.6837 | -105.4156
13.6874 | -105.425
13.6911 | -105.4356
13.6948 | -105.4493
13.6985 | -105.4602
13.7022 | -105.4693
13.7059 | -105.4818
13.7095 | -105.4926
13.7132 | -105.5018
13.7169 | -105.5122
13.7206 | -105.5256
13.7243 | -105.5381
13.728 | -105.5477
13.7317 | -105.5569
13.7354 | -105.5663
13.7391 | -105.5771
13.7427 | -105.5907
13.7464 | -105.6047
13.7501 | -105.6173
13.7538 | -105.6268
13.7575 | -105.6306
13.7612 | -105.6345
13.7649 | -105.6369
13.7686 | -105.639
13.7722 | -105.6401
13.7759 | -105.6412
13.7796 | -105.6413
13.7833 | -105.6407
13.787 | -105.6352
13.7907 | -105.6265
13.7944 | -105.6182
13.7981 | -105.6094
13.8018 | -105.5984
13.8054 | -105.5824
13.8091 | -105.5693
13.8128 | -105.5541
13.8165 | -105.5391
13.8202 | -105.5251
13.8239 | -105.5112
13.8276 | -105.4985
13.8313 | -105.4874
13.8349 | -105.4727
13.8386 | -105.4613
13.8423 | -105.4508
13.846 | -105.4392
13.8497 | -105.429
13.8534 | -105.4183
13.8571 | -105.4072
13.8608 | -105.4004
13.8645 | -105.3932
13.8681 | -105.388
13.8718 | -105.3849
13.8755 | -105.3823
13.8792 | -105.3802
13.8829 | -105.3769
13.8866 | -105.3779
13.8903 | -105.3801
13.894 | -105.3804
13.8976 | -105.3819
13.9013 | -105.3843
13.905 | -105.3827
13.9087 | -105.384
13.9124 | -105.3881
13.9161 | -105.3903
13.9198 | -105.3906
13.9235 | -105.3937
13.9272 | -105.4014
13.9308 | -105.4073
13.9345 | -105.4124
13.9382 | -105.4176
13.9419 | -105.4238
13.9456 | -105.4295
13.9493 | -105.4396
13.953 | -105.4512
13.9567 | -105.4608
13.9604 | -105.4668
13.964 | -105.4681
13.9677 | -105.4712
13.9714 | -105.4773
13.9751 | -105.4826
13.9788 | -105.489
13.9825 | -105.4928
13.9862 | -105.4933
13.9899 | -105.4945
13.9935 | -105.4939
13.9972 | -105.4937
14.0009 | -105.4886
14.0046 | -105.4832
14.0083 | -105.4723
14.012 | -105.4611
14.0157 | -105.4499
14.0194 | -105.4365
14.0231 | -105.4187
14.0267 | -105.3998
14.0304 | -105.3802
14.0341 | -105.3607
14.0378 | -105.3427
14.0415 | -105.3237
14.0452 | -105.3032
14.0489 | -105.2828
14.0526 | -105.2638
14.0562 | -105.247
14.0599 | -105.233
14.0636 | -105.2203
14.0673 | -105.2151
14.071 | -105.2125
14.0747 | -105.2124
14.0784 | -105.216
14.0821 | -105.2215
14.0858 | -105.2312
14.0894 | -105.2438
14.0931 | -105.2586
14.0968 | -105.2747
14.1005 | -105.2935
14.1042 | -105.3104
14.1079 | -105.329
14.1116 | -105.347
14.1153 | -105.3643
14.1189 | -105.3832
14.1226 | -105.4006
14.1263 | -105.4152
14.13 | -105.4258
14.1337 | -105.4361
14.1374 | -105.4452
14.1411 | -105.4503
14.1448 | -105.4525
14.1485 | -105.4541
14.1521 | -105.4509
14.1558 | -105.4449
14.1595 | -105.4393
14.1632 | -105.436
14.1669 | -105.4319
14.1706 | -105.4232
14.1743 | -105.4103
14.178 | -105.3954
14.1817 | -105.3801
14.1853 | -105.3654
14.189 | -105.3548
14.1927 | -105.3443
14.1964 | -105.3346
14.2001 | -105.3258
14.2038 | -105.3175
14.2075 | -105.3107
14.2112 | -105.3058
14.2148 | -105.3036
14.2185 | -105.3016
14.2222 | -105.3006
14.2259 | -105.3041
14.2296 | -105.3075
14.2333 | -105.3144
14.237 | -105.3252
14.2407 | -105.3349
14.2444 | -105.3456
14.248 | -105.3575
14.2517 | -105.3685
14.2554 | -105.3811
14.2591 | -105.3943
14.2628 | -105.4044
14.2665 | -105.412
14.2702 | -105.4219
14.2739 | -105.4337
14.2775 | -105.4439
14.2812 | -105.4552
14.2849 | -105.4655
14.2886 | -105.4738
14.2923 | -105.4818
14.296 | -105.4896
14.2997 | -105.4995
14.3034 | -105.5091
14.3071 | -105.5189
14.3107 | -105.5286
14.3144 | -105.5343
14.3181 | -105.5387
14.3218 | -105.5444
14.3255 | -105.5502
14.3292 | -105.5512
14.3329 | -105.5515
14.3366 | -105.5538
14.3402 | -105.5556
14.3439 | -105.5585
14.3476 | -105.5614
14.3513 | -105.5661
14.355 | -105.5711
14.3587 | -105.5742
14.3624 | -105.578
14.3661 | -105.582
14.3698 | -105.586
14.3734 | -105.5908
14.3771 | -105.5917
14.3808 | -105.5912
14.3845 | -105.5936
14.3882 | -105.5993
14.3919 | -105.6034
14.3956 | -105.606
14.3993 | -105.6059
14.403 | -105.6092
14.4066 | -105.6126
14.4103 | -105.6168
14.414 | -105.6186
14.4177 | -105.6216
14.4214 | -105.6233
14.4251 | -105.6265
14.4288 | -105.6279
14.4325 | -105.6259
14.4361 | -105.6254
14.4398 | -105.6248
14.4435 | -105.6234
14.4472 | -105.6234
14.4509 | -105.6236
14.4546 | -105.6253
14.4583 | -105.6277
14.462 | -105.6256
14.4657 | -105.6241
14.4693 | -105.6233
14.473 | -105.6249
14.4767 | -105.623
14.4804 | -105.6203
14.4841 | -105.6175
14.4878 | -105.6145
14.4915 | -105.6139
14.4952 | -105.6135
14.4988 | -105.6121
14.5025 | -105.6114
14.5062 | -105.6098
14.5099 | -105.6105
14.5136 | -105.6123
14.5173 | -105.6143
14.521 | -105.6163
14.5247 | -105.6192
14.5284 | -105.6196
14.532 | -105.6199
14.5357 | -105.6218
14.5394 | -105.6235
14.5431 | -105.6246
14.5468 | -105.6243
14.5505 | -105.6216
14.5542 | -105.6197
14.5579 | -105.6171
14.5615 | -105.6157
14.5652 | -105.6157
14.5689 | -105.6156
14.5726 | -105.6155
14.5763 | -105.6137
14.58 | -105.6104
14.5837 | -105.6055
14.5874 | -105.6043
14.5911 | -105.606
14.5947 | -105.6044
14.5984 | -105.5962
14.6021 | -105.5937
14.6058 | -105.5927
14.6095 | -105.5936
14.6132 | -105.5903
14.6169 | -105.5873
14.6206 | -105.5858
14.6243 | -105.5858
14.6279 | -105.5853
14.6316 | -105.5859
14.6353 | -105.5859
14.639 | -105.5865
14.6427 | -105.589
14.6464 | -105.5908
14.6501 | -105.5925
14.6538 | -105.5941
14.6574 | -105.5962
14.6611 | -105.5972
14.6648 | -105.5973
14.6685 | -105.5993
14.6722 | -105.605
14.6759 | -105.6109
14.6796 | -105.6183
14.6833 | -105.625
14.687 | -105.6322
14.6906 | -105.6389
14.6943 | -105.6443
14.698 | -105.652
14.7017 | -105.6572
14.7054 | -105.6614
14.7091 | -105.6666
14.7128 | -105.6709
14.7165 | -105.6757
14.7201 | -105.6773
14.7238 | -105.6779
14.7275 | -105.6787
14.7312 | -105.6792
14.7349 | -105.6762
14.7386 | -105.6742
14.7423 | -105.6723
14.746 | -105.6687
14.7497 | -105.6634
14.7533 | -105.657
14.757 | -105.6499
14.7607 | -105.6457
14.7644 | -105.64
14.7681 | -105.6326
14.7718 | -105.6264
14.7755 | -105.6217
14.7792 | -105.6193
14.7828 | -105.6178
14.7865 | -105.6203
14.7902 | -105.6251
14.7939 | -105.6347
14.7976 | -105.6426
14.8013 | -105.65
14.805 | -105.6592
14.8087 | -105.6761
14.8124 | -105.692
14.816 | -105.7073
14.8197 | -105.7203
14.8234 | -105.7354
14.8271 | -105.7531
14.8308 | -105.7748
14.8345 | -105.7973
14.8382 | -105.8191
14.8419 | -105.8369
14.8456 | -105.8566
14.8492 | -105.8766
14.8529 | -105.8941
14.8566 | -105.9122
14.8603 | -105.9282
14.864 | -105.9415
14.8677 | -105.9519
14.8714 | -105.9608
14.8751 | -105.964
14.8787 | -105.9628
14.8824 | -105.9588
14.8861 | -105.9495
14.8898 | -105.9383
14.8935 | -105.9253
14.8972 | -105.8927
14.9009 | -105.8707
14.9046 | -105.8515
14.9083 | -105.8318
14.9119 | -105.8051
14.9156 | -105.7769
14.9193 | -105.7489
14.923 | -105.7193
14.9267 | -105.6894
14.9304 | -105.6604
14.9341 | -105.6339
14.9378 | -105.6099
14.9414 | -105.5871
14.9451 | -105.5649
14.9488 | -105.543
14.9525 | -105.5225
14.9562 | -105.5038
14.9599 | -105.4867
14.9636 | -105.4718
14.9673 | -105.4595
14.971 | -105.45
14.9746 | -105.4433
14.9783 | -105.4386
14.982 | -105.4374
14.9857 | -105.437
14.9894 | -105.4391
14.9931 | -105.4421
14.9968 | -105.4477
15.0005 | -105.4544
15.0041 | -105.4624
15.0078 | -105.4719
15.0115 | -105.4851
15.0152 | -105.4982
15.0189 | -105.5123
15.0226 | -105.5265
15.0263 | -105.541
15.03 | -105.5564
15.0337 | -105.5768
15.0373 | -105.5952
15.041 | -105.6108
15.0447 | -105.623
15.0484 | -105.6364
15.0521 | -105.6511
15.0558 | -105.6687
15.0595 | -105.683
15.0632 | -105.6974
15.0669 | -105.7116
15.0705 | -105.7214
15.0742 | -105.732
15.0779 | -105.7468
15.0816 | -105.7624
15.0853 | -105.778
15.089 | -105.7965
15.0927 | -105.8136
15.0964 | -105.8254
15.1 | -105.8383
15.1037 | -105.8533
15.1074 | -105.8672
15.1111 | -105.8816
15.1148 | -105.8977
15.1185 | -105.9149
15.1222 | -105.9283
15.1259 | -105.9401
15.1296 | -105.9519
15.1332 | -105.9624
15.1369 | -105.969
15.1406 | -105.9732
15.1443 | -105.9778
15.148 | -105.9822
15.1517 | -105.9849
15.1554 | -105.9852
15.1591 | -105.9819
15.1627 | -105.977
15.1664 | -105.9725
15.1701 | -105.9671
15.1738 | -105.9623
15.1775 | -105.9556
15.1812 | -105.9466
15.1849 | -105.9357
15.1886 | -105.9223
15.1923 | -105.9021
15.1959 | -105.8824
15.1996 | -105.8624
15.2033 | -105.8453
15.207 | -105.826
15.2107 | -105.8043
15.2144 | -105.7811
15.2181 | -105.7594
15.2218 | -105.7377
15.2255 | -105.716
15.2291 | -105.6959
15.2328 | -105.6767
15.2365 | -105.6588
15.2402 | -105.6396
15.2439 | -105.6207
15.2476 | -105.6038
15.2513 | -105.5865
15.255 | -105.5744
15.2586 | -105.5629
15.2623 | -105.5535
15.266 | -105.5464
15.2697 | -105.5407
15.2734 | -105.5356
15.2771 | -105.5337
15.2808 | -105.5346
15.2845 | -105.5371
15.2882 | -105.5437
15.2918 | -105.5504
15.2955 | -105.5546
15.2992 | -105.5595
15.3029 | -105.5666
15.3066 | -105.5752
15.3103 | -105.5886
15.314 | -105.6061
15.3177 | -105.6201
15.3213 | -105.6334
15.325 | -105.6475
15.3287 | -105.6636
15.3324 | -105.679
15.3361 | -105.6951
15.3398 | -105.7092
15.3435 | -105.7251
15.3472 | -105.7435
15.3509 | -105.7617
15.3545 | -105.7802
15.3582 | -105.7942
15.3619 | -105.8053
15.3656 | -105.8159
15.3693 | -105.8266
15.373 | -105.835
15.3767 | -105.8459
15.3804 | -105.8573
15.384 | -105.866
15.3877 | -105.8728
15.3914 | -105.8777
15.3951 | -105.8833
15.3988 | -105.8856
15.4025 | -105.8862
15.4062 | -105.89
15.4099 | -105.8956
15.4136 | -105.899
15.4172 | -105.8981
15.4209 | -105.8921
15.4246 | -105.8863
15.4283 | -105.8884
15.432 | -105.8891
15.4357 | -105.8885
15.4394 | -105.8873
15.4431 | -105.887
15.4467 | -105.8838
15.4504 | -105.8771
15.4541 | -105.8745
15.4578 | -105.8728
15.4615 | -105.8695
15.4652 | -105.8683
15.4689 | -105.8686
15.4726 | -105.8698
15.4763 | -105.8694
15.4799 | -105.8703
15.4836 | -105.8713
15.4873 | -105.8697
15.491 | -105.8692
15.4947 | -105.8716
15.4984 | -105.8739
15.5021 | -105.8751
15.5058 | -105.8784
15.5095 | -105.8832
15.5131 | -105.8877
15.5168 | -105.8927
15.5205 | -105.8991
15.5242 | -105.9055
15.5279 | -105.9103
15.5316 | -105.9116
15.5353 | -105.9137
15.539 | -105.923
15.5426 | -105.9314
15.5463 | -105.9377
15.55 | -105.9406
15.5537 | -105.9463
15.5574 | -105.9532
15.5611 | -105.9606
15.5648 | -105.9674
15.5685 | -105.9734
15.5722 | -105.9784
15.5758 | -105.9836
15.5795 | -105.9916
15.5832 | -105.999
15.5869 | -106.0021
15.5906 | -106.0042
15.5943 | -106.0098
15.598 | -106.0128
15.6017 | -106.0146
15.6053 | -106.0164
15.609 | -106.017
15.6127 | -106.0155
15.6164 | -106.0141
15.6201 | -106.0136
15.6238 | -106.0131
15.6275 | -106.0108
15.6312 | -106.0091
15.6349 | -106.0071
15.6385 | -106.0066
15.6422 | -106.0075
15.6459 | -106.0078
15.6496 | -106.0011
15.6533 | -105.995
15.657 | -105.9898
15.6607 | -105.9853
15.6644 | -105.9818
15.668 | -105.9761
15.6717 | -105.9727
15.6754 | -105.9709
15.6791 | -105.9661
15.6828 | -105.961
15.6865 | -105.9592
15.6902 | -105.953
15.6939 | -105.9437
15.6976 | -105.9334
15.7012 | -105.9194
15.7049 | -105.9064
15.7086 | -105.893
15.7123 | -105.8803
15.716 | -105.8675
15.7197 | -105.851
15.7234 | -105.8372
15.7271 | -105.8228
15.7308 | -105.8033
15.7344 | -105.7862
15.7381 | -105.7726
15.7418 | -105.7578
15.7455 | -105.7417
15.7492 | -105.7244
15.7529 | -105.7082
15.7566 | -105.6913
15.7603 | -105.6764
15.7639 | -105.6595
15.7676 | -105.643
15.7713 | -105.6292
15.775 | -105.6178
15.7787 | -105.6051
15.7824 | -105.5932
15.7861 | -105.5829
15.7898 | -105.5703
15.7935 | -105.56
15.7971 | -105.5502
15.8008 | -105.5406
15.8045 | -105.5317
15.8082 | -105.5226
15.8119 | -105.5159
15.8156 | -105.5097
15.8193 | -105.5031
15.823 | -105.4972
15.8266 | -105.494
15.8303 | -105.4946
15.834 | -105.4959
15.8377 | -105.498
15.8414 | -105.5015
15.8451 | -105.5073
15.8488 | -105.5123
15.8525 | -105.52
15.8562 | -105.527
15.8598 | -105.5333
15.8635 | -105.5412
15.8672 | -105.5491
15.8709 | -105.5573
15.8746 | -105.5672
15.8783 | -105.5793
15.882 | -105.5909
15.8857 | -105.6024
15.8894 | -105.6139
15.893 | -105.6251
15.8967 | -105.6363
15.9004 | -105.649
15.9041 | -105.6591
15.9078 | -105.6672
15.9115 | -105.676
15.9152 | -105.6878
15.9189 | -105.7009
15.9225 | -105.7135
15.9262 | -105.7266
15.9299 | -105.738
15.9336 | -105.7492
15.9373 | -105.7589
15.941 | -105.7648
15.9447 | -105.7724
15.9484 | -105.7786
15.9521 | -105.7826
15.9557 | -105.79
15.9594 | -105.7922
15.9631 | -105.7894
15.9668 | -105.7878
15.9705 | -105.7894
15.9742 | -105.7912
15.9779 | -105.7921
15.9816 | -105.7919
15.9852 | -105.7912
15.9889 | -105.7871
15.9926 | -105.7818
15.9963 | -105.7763
16 | -105.7701
16.0037 | -105.7631
16.0074 | -105.754
16.0111 | -105.7424
16.0148 | -105.7333
16.0184 | -105.7239
16.0221 | -105.7112
16.0258 | -105.6964
16.0295 | -105.6835
16.0332 | -105.6721
16.0369 | -105.6619
16.0406 | -105.6506
16.0443 | -105.638
16.0479 | -105.6256
16.0516 | -105.6141
16.0553 | -105.6017
16.059 | -105.5872
16.0627 | -105.5752
16.0664 | -105.5644
16.0701 | -105.5538
16.0738 | -105.5446
16.0775 | -105.5385
16.0811 | -105.5321
16.0848 | -105.5273
16.0885 | -105.5259
16.0922 | -105.5239
16.0959 | -105.5246
16.0996 | -105.5279
16.1033 | -105.5301
16.107 | -105.5297
16.1106 | -105.5319
16.1143 | -105.5379
16.118 | -105.543
16.1217 | -105.5488
16.1254 | -105.5537
16.1291 | -105.5587
16.1328 | -105.5683
16.1365 | -105.5776
16.1402 | -105.5847
16.1438 | -105.5924
16.1475 | -105.6036
16.1512 | -105.6164
16.1549 | -105.6302
16.1586 | -105.6458
16.1623 | -105.6618
16.166 | -105.6782
16.1697 | -105.6954
16.1734 | -105.7113
16.177 | -105.7277
16.1807 | -105.7457
16.1844 | -105.7615
16.1881 | -105.7781
16.1918 | -105.7913
16.1955 | -105.8065
16.1992 | -105.8217
16.2029 | -105.8361
16.2065 | -105.8501
16.2102 | -105.8651
16.2139 | -105.8806
16.2176 | -105.8972
16.2213 | -105.9115
16.225 | -105.9225
16.2287 | -105.9327
16.2324 | -105.9442
16.2361 | -105.9539
16.2397 | -105.9621
16.2434 | -105.9718
16.2471 | -105.9801
16.2508 | -105.9905
16.2545 | -106.0009
16.2582 | -106.0124
16.2619 | -106.0212
16.2656 | -106.0274
16.2692 | -106.0311
16.2729 | -106.0358
16.2766 | -106.0435
16.2803 | -106.0495
16.284 | -106.0551
16.2877 | -106.0599
16.2914 | -106.0636
16.2951 | -106.0641
16.2988 | -106.0677
16.3024 | -106.0712
16.3061 | -106.0727
16.3098 | -106.0731
16.3135 | -106.0729
16.3172 | -106.0747
16.3209 | -106.0759
16.3246 | -106.0771
16.3283 | -106.0773
16.3319 | -106.0761
16.3356 | -106.0761
16.3393 | -106.0778
16.343 | -106.0788
16.3467 | -106.0771
16.3504 | -106.0783
16.3541 | -106.0811
16.3578 | -106.0841
16.3615 | -106.0846
16.3651 | -106.0843
16.3688 | -106.086
16.3725 | -106.0862
16.3762 | -106.0865
16.3799 | -106.0855
16.3836 | -106.0838
16.3873 | -106.0846
16.391 | -106.0823
16.3947 | -106.0806
16.3983 | -106.0826
16.402 | -106.0876
16.4057 | -106.0927
16.4094 | -106.0974
16.4131 | -106.1029
16.4168 | -106.1085
16.4205 | -106.1137
16.4242 | -106.1243
16.4278 | -106.132
16.4315 | -106.1399
16.4352 | -106.1464
16.4389 | -106.1527
16.4426 | -106.1584
16.4463 | -106.1641
16.45 | -106.1691
16.4537 | -106.1745
16.4574 | -106.1842
16.461 | -106.1952
16.4647 | -106.2007
16.4684 | -106.2074
16.4721 | -106.2165
16.4758 | -106.2204
16.4795 | -106.2222
16.4832 | -106.2236
16.4869 | -106.2249
16.4905 | -106.2273
16.4942 | -106.2234
16.4979 | -106.2219
16.5016 | -106.2189
16.5053 | -106.2166
16.509 | -106.2133
16.5127 | -106.2054
16.5164 | -106.1974
16.5201 | -106.1907
16.5237 | -106.1836
16.5274 | -106.1772
16.5311 | -106.1719
16.5348 | -106.1637
16.5385 | -106.1548
16.5422 | -106.1464
16.5459 | -106.1383
16.5496 | -106.1299
16.5532 | -106.1195
16.5569 | -106.1069
16.5606 | -106.0963
16.5643 | -106.0857
16.568 | -106.0771
16.5717 | -106.0682
16.5754 | -106.06
16.5791 | -106.0497
16.5828 | -106.0411
16.5864 | -106.0297
16.5901 | -106.0192
16.5938 | -106.0093
16.5975 | -106.0024
16.6012 | -105.9963
16.6049 | -105.9909
16.6086 | -105.9869
16.6123 | -105.982
16.616 | -105.9756
16.6196 | -105.9727
16.6233 | -105.9711
16.627 | -105.9735
16.6307 | -105.9759
16.6344 | -105.9767
16.6381 | -105.9763
16.6418 | -105.9773
16.6455 | -105.9794
16.6491 | -105.9853
16.6528 | -105.9922
16.6565 | -106.0006
16.6602 | -106.0097
16.6639 | -106.0196
16.6676 | -106.0299
16.6713 | -106.0419
16.675 | -106.0553
16.6787 | -106.0693
16.6823 | -106.0834
16.686 | -106.0956
16.6897 | -106.1086
16.6934 | -106.1214
16.6971 | -106.1346
16.7008 | -106.1504
16.7045 | -106.1629
16.7082 | -106.1735
16.7118 | -106.1863
16.7155 | -106.1983
16.7192 | -106.2088
16.7229 | -106.2168
16.7266 | -106.2241
16.7303 | -106.2314
16.734 | -106.2383
16.7377 | -106.2457
16.7414 | -106.2504
16.745 | -106.2542
16.7487 | -106.2546
16.7524 | -106.2563
16.7561 | -106.257
16.7598 | -106.2528
16.7635 | -106.2496
16.7672 | -106.2441
16.7709 | -106.2371
16.7746 | -106.2306
16.7782 | -106.2242
16.7819 | -106.217
16.7856 | -106.2081
16.7893 | -106.1988
16.793 | -106.1899
16.7967 | -106.1839
16.8004 | -106.1757
16.8041 | -106.1657
16.8077 | -106.1543
16.8114 | -106.1453
16.8151 | -106.1381
16.8188 | -106.1328
16.8225 | -106.1254
16.8262 | -106.1217
16.8299 | -106.1178
16.8336 | -106.1138
16.8373 | -106.109
16.8409 | -106.1032
16.8446 | -106.0991
16.8483 | -106.0942
16.852 | -106.0894
16.8557 | -106.0862
16.8594 | -106.0835
16.8631 | -106.0813
16.8668 | -106.0783
16.8704 | -106.0744
16.8741 | -106.0695
16.8778 | -106.0682
16.8815 | -106.0687
16.8852 | -106.0661
16.8889 | -106.0595
16.8926 | -106.0551
16.8963 | -106.0512
16.9 | -106.0449
16.9036 | -106.0381
16.9073 | -106.0317
16.911 | -106.0256
16.9147 | -106.0215
16.9184 | -106.0163
16.9221 | -106.0116
16.9258 | -106.0084
16.9295 | -106.0026
16.9331 | -105.9957
16.9368 | -105.9896
16.9405 | -105.9819
16.9442 | -105.9772
16.9479 | -105.9721
16.9516 | -105.9676
16.9553 | -105.9615
16.959 | -105.9543
16.9627 | -105.9489
16.9663 | -105.9437
16.97 | -105.9386
16.9737 | -105.9335
16.9774 | -105.9283
16.9811 | -105.9258
16.9848 | -105.9244
16.9885 | -105.924
16.9922 | -105.9236
16.9959 | -105.9224
16.9995 | -105.9228
17.0032 | -105.9243
17.0069 | -105.927
17.0106 | -105.9302
17.0143 | -105.9326
17.018 | -105.9354
17.0217 | -105.9414
17.0254 | -105.9498
17.029 | -105.9564
17.0327 | -105.9646
17.0364 | -105.9732
17.0401 | -105.9805
17.0438 | -105.9879
17.0475 | -105.9957
17.0512 | -106.0036
17.0549 | -106.0122
17.0586 | -106.02
17.0622 | -106.0276
17.0659 | -106.0349
17.0696 | -106.0392
17.0733 | -106.0452
17.077 | -106.0534
17.0807 | -106.0595
17.0844 | -106.0658
17.0881 | -106.0709
17.0917 | -106.0768
17.0954 | -106.0834
17.0991 | -106.0895
17.1028 | -106.0924
17.1065 | -106.0913
17.1102 | -106.091
17.1139 | -106.0921
17.1176 | -106.0923
17.1213 | -106.0908
17.1249 | -106.0887
17.1286 | -106.0884
17.1323 | -106.0893
17.136 | -106.0875
17.1397 | -106.0844
17.1434 | -106.0831
17.1471 | -106.0817
17.1508 | -106.078
17.1544 | -106.0745
17.1581 | -106.0679
17.1618 | -106.0623
17.1655 | -106.0564
17.1692 | -106.0501
17.1729 | -106.0458
17.1766 | -106.0426
17.1803 | -106.0371
17.184 | -106.0303
17.1876 | -106.0213
17.1913 | -106.0156
17.195 | -106.0069
17.1987 | -106.0013
17.2024 | -106.0002
17.2061 | -105.9981
17.2098 | -105.9933
17.2135 | -105.9864
17.2172 | -105.9777
17.2208 | -105.9709
17.2245 | -105.9646
17.2282 | -105.9586
17.2319 | -105.9534
17.2356 | -105.9479
17.2393 | -105.9427
17.243 | -105.9387
17.2467 | -105.9339
17.2503 | -105.9289
17.254 | -105.9258
17.2577 | -105.924
17.2614 | -105.9199
17.2651 | -105.9147
17.2688 | -105.9112
17.2725 | -105.9098
17.2762 | -105.9059
17.2799 | -105.9015
17.2835 | -105.8977
17.2872 | -105.8962
17.2909 | -105.8938
17.2946 | -105.8927
17.2983 | -105.8922
17.302 | -105.8911
17.3057 | -105.889
17.3094 | -105.8854
17.313 | -105.8823
17.3167 | -105.8815
17.3204 | -105.8809
17.3241 | -105.8786
17.3278 | -105.8782
17.3315 | -105.8797
17.3352 | -105.8823
17.3389 | -105.8832
17.3426 | -105.8818
17.3462 | -105.8819
17.3499 | -105.8823
17.3536 | -105.8852
17.3573 | -105.8882
17.361 | -105.8924
17.3647 | -105.8987
17.3684 | -105.9043
17.3721 | -105.909
17.3757 | -105.915
17.3794 | -105.9229
17.3831 | -105.9315
17.3868 | -105.9401
17.3905 | -105.9494
17.3942 | -105.9614
17.3979 | -105.9764
17.4016 | -105.9922
17.4053 | -106.0088
17.4089 | -106.0278
17.4126 | -106.0488
17.4163 | -106.071
17.42 | -106.093
17.4237 | -106.1155
17.4274 | -106.1376
17.4311 | -106.1636
17.4348 | -106.1898
17.4385 | -106.2118
17.4421 | -106.2358
17.4458 | -106.2588
17.4495 | -106.2797
17.4532 | -106.2998
17.4569 | -106.3218
17.4606 | -106.3458
17.4643 | -106.3681
17.468 | -106.3851
17.4716 | -106.4024
17.4753 | -106.4194
17.479 | -106.434
17.4827 | -106.4463
17.4864 | -106.4568
17.4901 | -106.466
17.4938 | -106.4722
17.4975 | -106.4766
17.5012 | -106.4781
17.5048 | -106.4775
17.5085 | -106.479
17.5122 | -106.478
17.5159 | -106.4744
17.5196 | -106.4692
17.5233 | -106.4631
17.527 | -106.4555
17.5307 | -106.4467
17.5343 | -106.4339
17.538 | -106.4195
17.5417 | -106.4088
17.5454 | -106.3975
17.5491 | -106.3799
17.5528 | -106.3657
17.5565 | -106.3548
17.5602 | -106.3444
17.5639 | -106.3358
17.5675 | -106.3251
17.5712 | -106.3151
17.5749 | -106.3057
17.5786 | -106.295
17.5823 | -106.2826
17.586 | -106.2721
17.5897 | -106.2638
17.5934 | -106.2558
17.597 | -106.2478
17.6007 | -106.242
17.6044 | -106.2374
17.6081 | -106.2326
17.6118 | -106.227
17.6155 | -106.2214
17.6192 | -106.2163
17.6229 | -106.2113
17.6266 | -106.2068
17.6302 | -106.2032
17.6339 | -106.2009
17.6376 | -106.1989
17.6413 | -106.1953
17.645 | -106.1929
17.6487 | -106.1905
17.6524 | -106.1878
17.6561 | -106.1876
17.6598 | -106.1858
17.6634 | -106.184
17.6671 | -106.1841
17.6708 | -106.1849
17.6745 | -106.185
17.6782 | -106.1861
17.6819 | -106.1872
17.6856 | -106.1913
17.6893 | -106.1977
17.6929 | -106.2038
17.6966 | -106.2081
17.7003 | -106.2115
17.704 | -106.2157
17.7077 | -106.2216
17.7114 | -106.2265
17.7151 | -106.2335
17.7188 | -106.2422
17.7225 | -106.2478
17.7261 | -106.2523
17.7298 | -106.2583
17.7335 | -106.2642
17.7372 | -106.2716
17.7409 | -106.2755
17.7446 | -106.278
17.7483 | -106.2829
17.752 | -106.2862
17.7556 | -106.2895
17.7593 | -106.2918
17.763 | -106.2935
17.7667 | -106.294
17.7704 | -106.2929
17.7741 | -106.2907
17.7778 | -106.2898
17.7815 | -106.2867
17.7852 | -106.2841
17.7888 | -106.2836
17.7925 | -106.2779
17.7962 | -106.271
17.7999 | -106.2674
17.8036 | -106.2599
17.8073 | -106.2479
17.811 | -106.2396
17.8147 | -106.2319
17.8183 | -106.2228
17.822 | -106.2137
17.8257 | -106.2018
17.8294 | -106.1903
17.8331 | -106.1786
17.8368 | -106.1686
17.8405 | -106.1589
17.8442 | -106.1497
17.8479 | -106.1411
17.8515 | -106.1307
17.8552 | -106.1212
17.8589 | -106.1099
17.8626 | -106.1001
17.8663 | -106.092
17.87 | -106.084
17.8737 | -106.0756
17.8774 | -106.0712
17.8811 | -106.0681
17.8847 | -106.0639
17.8884 | -106.0603
17.8921 | -106.0595
17.8958 | -106.0623
17.8995 | -106.0634
17.9032 | -106.0637
17.9069 | -106.0657
17.9106 | -106.0689
17.9142 | -106.0745
17.9179 | -106.082
17.9216 | -106.0883
17.9253 | -106.0969
17.929 | -106.107
17.9327 | -106.1191
17.9364 | -106.131
17.9401 | -106.1419
17.9438 | -106.1488
17.9474 | -106.1573
17.9511 | -106.1679
17.9548 | -106.1753
17.9585 | -106.1822
17.9622 | -106.1873
17.9659 | -106.193
17.9696 | -106.1962
17.9733 | -106.1981
17.9769 | -106.2001
17.9806 | -106.2012
17.9843 | -106.1981
17.988 | -106.1938
17.9917 | -106.1888
17.9954 | -106.185
17.9991 | -106.1806
18.0028 | -106.1747
18.0065 | -106.1698
18.0101 | -106.1641
18.0138 | -106.1555
18.0175 | -106.1463
18.0212 | -106.1385
18.0249 | -106.1317
18.0286 | -106.1244
18.0323 | -106.1168
18.036 | -106.1111
18.0397 | -106.1065
18.0433 | -106.1018
18.047 | -106.0962
18.0507 | -106.0938
18.0544 | -106.0932
18.0581 | -106.0911
18.0618 | -106.0918
18.0655 | -106.0933
18.0692 | -106.0953
18.0728 | -106.0972
18.0765 | -106.1001
18.0802 | -106.1075
18.0839 | -106.115
18.0876 | -106.1214
18.0913 | -106.1293
18.095 | -106.1367
18.0987 | -106.1431
18.1024 | -106.1509
18.106 | -106.1604
18.1097 | -106.1697
18.1134 | -106.1768
18.1171 | -106.1878
18.1208 | -106.1978
18.1245 | -106.2066
18.1282 | -106.2152
18.1319 | -106.2236
18.1355 | -106.2297
18.1392 | -106.2346
18.1429 | -106.24
18.1466 | -106.2449
18.1503 | -106.2495
18.154 | -106.2509
18.1577 | -106.2507
18.1614 | -106.2503
18.1651 | -106.2512
18.1687 | -106.25
18.1724 | -106.2456
18.1761 | -106.239
18.1798 | -106.2337
18.1835 | -106.2287
18.1872 | -106.2226
18.1909 | -106.2147
18.1946 | -106.2055
18.1982 | -106.1955
18.2019 | -106.1881
18.2056 | -106.1804
18.2093 | -106.1718
18.213 | -106.1653
18.2167 | -106.1573
18.2204 | -106.1465
18.2241 | -106.1379
18.2278 | -106.1294
18.2314 | -106.1198
18.2351 | -106.1123
18.2388 | -106.1039
18.2425 | -106.0983
18.2462 | -106.0922
18.2499 | -106.0851
18.2536 | -106.0773
18.2573 | -106.0704
18.261 | -106.0606
18.2646 | -106.0524
18.2683 | -106.044
18.272 | -106.037
18.2757 | -106.0288
18.2794 | -106.02
18.2831 | -106.0121
18.2868 | -106.0046
18.2905 | -105.9963
18.2941 | -105.9891
18.2978 | -105.9826
18.3015 | -105.9735
18.3052 | -105.9645
18.3089 | -105.9553
18.3126 | -105.9481
18.3163 | -105.9417
18.32 | -105.934
18.3237 | -105.9247
18.3273 | -105.9146
18.331 | -105.9048
18.3347 | -105.8972
18.3384 | -105.8873
18.3421 | -105.877
18.3458 | -105.8711
18.3495 | -105.8662
18.3532 | -105.8568
18.3568 | -105.8453
18.3605 | -105.8352
18.3642 | -105.8266
18.3679 | -105.8188
18.3716 | -105.8125
18.3753 | -105.807
18.379 | -105.7992
18.3827 | -105.7923
18.3864 | -105.7884
18.39 | -105.7855
18.3937 | -105.7835
18.3974 | -105.7813
18.4011 | -105.779
18.4048 | -105.7774
18.4085 | -105.7764
18.4122 | -105.7766
18.4159 | -105.7767
18.4195 | -105.7794
18.4232 | -105.785
18.4269 | -105.7881
18.4306 | -105.7925
18.4343 | -105.7994
18.438 | -105.8048
18.4417 | -105.8105
18.4454 | -105.817
18.4491 | -105.8245
18.4527 | -105.8346
18.4564 | -105.8445
18.4601 | -105.8555
18.4638 | -105.8657
18.4675 | -105.8746
18.4712 | -105.886
18.4749 | -105.8977
18.4786 | -105.9101
18.4823 | -105.9247
18.4859 | -105.9367
18.4896 | -105.9507
18.4933 | -105.9647
18.497 | -105.9741
18.5007 | -105.9852
18.5044 | -105.9982
18.5081 | -106.011
18.5118 | -106.0211
18.5154 | -106.0314
18.5191 | -106.0403
18.5228 | -106.0504
18.5265 | -106.0566
18.5302 | -106.0625
18.5339 | -106.0689
18.5376 | -106.0772
18.5413 | -106.0856
18.545 | -106.0913
18.5486 | -106.0965
18.5523 | -106.099
18.556 | -106.0995
18.5597 | -106.1002
18.5634 | -106.1031
18.5671 | -106.108
18.5708 | -106.1078
18.5745 | -106.1063
18.5781 | -106.1041
18.5818 | -106.1043
18.5855 | -106.1028
18.5892 | -106.101
18.5929 | -106.0999
18.5966 | -106.095
18.6003 | -106.0915
18.604 | -106.0887
18.6077 | -106.0824
18.6113 | -106.0747
18.615 | -106.0697
18.6187 | -106.0658
18.6224 | -106.0566
18.6261 | -106.051
18.6298 | -106.0457
18.6335 | -106.0392
18.6372 | -106.033
18.6408 | -106.0277
18.6445 | -106.0218
18.6482 | -106.0155
18.6519 | -106.008
18.6556 | -106.0033
18.6593 | -105.9998
18.663 | -105.9951
18.6667 | -105.993
18.6704 | -105.9912
18.674 | -105.9902
18.6777 | -105.989
18.6814 | -105.9908
18.6851 | -105.992
18.6888 | -105.9906
18.6925 | -105.9888
18.6962 | -105.9883
18.6999 | -105.9898
18.7035 | -105.994
18.7072 | -105.9994
18.7109 | -106.004
18.7146 | -106.0087
18.7183 | -106.0117
18.722 | -106.0181
18.7257 | -106.0236
18.7294 | -106.0301
18.7331 | -106.0381
18.7367 | -106.045
18.7404 | -106.0519
18.7441 | -106.0592
18.7478 | -106.0666
18.7515 | -106.0742
18.7552 | -106.0794
18.7589 | -106.0868
18.7626 | -106.0977
18.7663 | -106.1066
18.7699 | -106.1137
18.7736 | -106.122
18.7773 | -106.1323
18.781 | -106.1423
18.7847 | -106.1521
18.7884 | -106.1605
18.7921 | -106.17
18.7958 | -106.1772
18.7994 | -106.1818
18.8031 | -106.1884
18.8068 | -106.193
18.8105 | -106.199
18.8142 | -106.2035
18.8179 | -106.2121
18.8216 | -106.2197
18.8253 | -106.2255
18.829 | -106.2309
18.8326 | -106.2344
18.8363 | -106.2394
18.84 | -106.2432
18.8437 | -106.2489
18.8474 | -106.2509
18.8511 | -106.2526
18.8548 | -106.2544
18.8585 | -106.2562
18.8621 | -106.2589
18.8658 | -106.2574
18.8695 | -106.2549
18.8732 | -106.2547
18.8769 | -106.2539
18.8806 | -106.2506
18.8843 | -106.2459
18.888 | -106.2412
18.8917 | -106.2362
18.8953 | -106.2295
18.899 | -106.2212
18.9027 | -106.2141
18.9064 | -106.2064
18.9101 | -106.1996
18.9138 | -106.1941
18.9175 | -106.1897
18.9212 | -106.1841
18.9248 | -106.1776
18.9285 | -106.1709
18.9322 | -106.1657
18.9359 | -106.1608
18.9396 | -106.157
18.9433 | -106.1504
18.947 | -106.1447
18.9507 | -106.1399
18.9544 | -106.1349
18.958 | -106.1316
18.9617 | -106.1311
18.9654 | -106.1314
18.9691 | -106.1304
18.9728 | -106.1292
18.9765 | -106.1289
18.9802 | -106.1278
18.9839 | -106.1273
18.9876 | -106.1298
18.9912 | -106.1314
18.9949 | -106.1342
18.9986 | -106.1368
19.0023 | -106.1403
19.006 | -106.1416
19.0097 | -106.1411
19.0134 | -106.1429
19.0171 | -106.1455
19.0207 | -106.1484
19.0244 | -106.153
19.0281 | -106.1565
19.0318 | -106.1578
19.0355 | -106.1605
19.0392 | -106.1626
19.0429 | -106.1644
19.0466 | -106.1643
19.0503 | -106.164
19.0539 | -106.1658
19.0576 | -106.1677
19.0613 | -106.1687
19.065 | -106.1721
19.0687 | -106.1749
19.0724 | -106.1721
19.0761 | -106.1724
19.0798 | -106.1714
19.0834 | -106.1722
19.0871 | -106.1696
19.0908 | -106.1685
19.0945 | -106.1664
19.0982 | -106.1636
19.1019 | -106.1606
19.1056 | -106.1586
19.1093 | -106.156
19.113 | -106.1518
19.1166 | -106.1469
19.1203 | -106.1409
19.124 | -106.1344
19.1277 | -106.1275
19.1314 | -106.1201
19.1351 | -106.1143
19.1388 | -106.1088
19.1425 | -106.0994
19.1461 | -106.0896
19.1498 | -106.0787
19.1535 | -106.0641
19.1572 | -106.0544
19.1609 | -106.0435
19.1646 | -106.0346
19.1683 | -106.0242
19.172 | -106.0111
19.1757 | -105.9983
19.1793 | -105.9845
19.183 | -105.9711
19.1867 | -105.9588
19.1904 | -105.9463
19.1941 | -105.9349
19.1978 | -105.9226
19.2015 | -105.9113
19.2052 | -105.9035
19.2089 | -105.8943
19.2125 | -105.8842
19.2162 | -105.8796
19.2199 | -105.8746
19.2236 | -105.8713
19.2273 | -105.8694
19.231 | -105.8657
19.2347 | -105.8631
19.2384 | -105.86
19.242 | -105.859
19.2457 | -105.8606
19.2494 | -105.8611
19.2531 | -105.8653
19.2568 | -105.8698
19.2605 | -105.8712
19.2642 | -105.874
19.2679 | -105.8768
19.2716 | -105.8823
19.2752 | -105.8872
19.2789 | -105.8942
19.2826 | -105.9025
19.2863 | -105.9117
19.29 | -105.9189
19.2937 | -105.9255
19.2974 | -105.9333
19.3011 | -105.9418
19.3047 | -105.9493
19.3084 | -105.9573
19.3121 | -105.9666
19.3158 | -105.9755
19.3195 | -105.9848
19.3232 | -105.9927
19.3269 | -105.9984
19.3306 | -106.0042
19.3343 | -106.0096
19.3379 | -106.0143
19.3416 | -106.018
19.3453 | -106.0224
19.349 | -106.0245
19.3527 | -106.0264
19.3564 | -106.0263
19.3601 | -106.023
19.3638 | -106.0166
19.3675 | -106.0129
19.3711 | -106.0096
19.3748 | -106.007
19.3785 | -106.0031
19.3822 | -105.9964
19.3859 | -105.9906
19.3896 | -105.9839
19.3933 | -105.9775
19.397 | -105.9713
19.4006 | -105.9645
19.4043 | -105.9561
19.408 | -105.946
19.4117 | -105.9358
19.4154 | -105.9288
19.4191 | -105.9208
19.4228 | -105.9139
19.4265 | -105.9056
19.4302 | -105.896
19.4338 | -105.8876
19.4375 | -105.8789
19.4412 | -105.8692
19.4449 | -105.8611
19.4486 | -105.8532
19.4523 | -105.8454
19.456 | -105.839
19.4597 | -105.8288
19.4633 | -105.8171
19.467 | -105.8058
19.4707 | -105.7968
19.4744 | -105.7871
19.4781 | -105.777
19.4818 | -105.7649
19.4855 | -105.7546
19.4892 | -105.7464
19.4929 | -105.7361
19.4965 | -105.728
19.5002 | -105.7215
19.5039 | -105.7157
19.5076 | -105.7113
19.5113 | -105.707
19.515 | -105.7033
19.5187 | -105.6999
19.5224 | -105.6968
19.526 | -105.6955
19.5297 | -105.6937
19.5334 | -105.6951
19.5371 | -105.6986
19.5408 | -105.7017
19.5445 | -105.7063
19.5482 | -105.7064
19.5519 | -105.7078
19.5556 | -105.7133
19.5592 | -105.719
19.5629 | -105.7231
19.5666 | -105.7285
19.5703 | -105.7361
19.574 | -105.7426
19.5777 | -105.7498
19.5814 | -105.759
19.5851 | -105.7655
19.5888 | -105.7712
19.5924 | -105.7785
19.5961 | -105.7862
19.5998 | -105.7938
19.6035 | -105.8021
19.6072 | -105.8085
19.6109 | -105.8147
19.6146 | -105.8208
19.6183 | -105.8281
19.6219 | -105.8336
19.6256 | -105.8358
19.6293 | -105.8384
19.633 | -105.8413
19.6367 | -105.8472
19.6404 | -105.8523
19.6441 | -105.8554
19.6478 | -105.859
19.6515 | -105.8591
19.6551 | -105.8575
19.6588 | -105.8593
19.6625 | -105.8607
19.6662 | -105.8612
19.6699 | -105.8614
19.6736 | -105.8615
19.6773 | -105.861
19.681 | -105.8615
19.6846 | -105.8629
19.6883 | -105.8651
19.692 | -105.8638
19.6957 | -105.8611
19.6994 | -105.8591
19.7031 | -105.8569
19.7068 | -105.8562
19.7105 | -105.8556
19.7142 | -105.8526
19.7178 | -105.8515
19.7215 | -105.851
19.7252 | -105.8512
19.7289 | -105.8543
19.7326 | -105.8562
19.7363 | -105.8574
19.74 | -105.8607
19.7437 | -105.8654
19.7473 | -105.87
19.751 | -105.8714
19.7547 | -105.8742
19.7584 | -105.8791
19.7621 | -105.8838
19.7658 | -105.8865
19.7695 | -105.8902
19.7732 | -105.8941
19.7769 | -105.8979
19.7805 | -105.8996
19.7842 | -105.9055
19.7879 | -105.913
19.7916 | -105.9197
19.7953 | -105.9277
19.799 | -105.9355
19.8027 | -105.9429
19.8064 | -105.949
19.8101 | -105.9549
19.8137 | -105.9587
19.8174 | -105.9649
19.8211 | -105.9713
19.8248 | -105.9756
19.8285 | -105.9837
19.8322 | -105.9904
19.8359 | -105.9922
19.8396 | -105.9952
19.8432 | -105.9972
19.8469 | -105.9994
19.8506 | -106
19.8543 | -106.0006
19.858 | -106.0015
19.8617 | -106.0028
19.8654 | -106.0036
19.8691 | -106.0035
19.8728 | -106.0032
19.8764 | -106.0027
19.8801 | -106.0034
19.8838 | -106.0037
19.8875 | -106.0034
19.8912 | -106.0039
19.8949 | -106.004
19.8986 | -106.0041
19.9023 | -106.0043
19.9059 | -106.0077
19.9096 | -106.0099
19.9133 | -106.011
19.917 | -106.0105
19.9207 | -106.013
19.9244 | -106.016
19.9281 | -106.0179
19.9318 | -106.0217
19.9355 | -106.0282
19.9391 | -106.0344
19.9428 | -106.0401
19.9465 | -106.0485
19.9502 | -106.0579
19.9539 | -106.0657
19.9576 | -106.0735
19.9613 | -106.0828
19.965 | -106.0918
19.9686 | -106.1017
19.9723 | -106.1126
19.976 | -106.1278
19.9797 | -106.1429
19.9834 | -106.1558
19.9871 | -106.1692
19.9908 | -106.1816
19.9945 | -106.1947
19.9982 | -106.2077
20.0018 | -106.2171
20.0055 | -106.2264
20.0092 | -106.2361
20.0129 | -106.2451
20.0166 | -106.2558
20.0203 | -106.2674
20.024 | -106.2785
20.0277 | -106.2873
20.0314 | -106.2949
20.035 | -106.3036
20.0387 | -106.3121
20.0424 | -106.3179
20.0461 | -106.3217
20.0498 | -106.3239
20.0535 | -106.3265
20.0572 | -106.3287
20.0609 | -106.3279
20.0645 | -106.3273
20.0682 | -106.3223
20.0719 | -106.316
20.0756 | -106.3084
20.0793 | -106.3008
20.083 | -106.2914
20.0867 | -106.2813
20.0904 | -106.2737
20.0941 | -106.2653
20.0977 | -106.2585
20.1014 | -106.2483
20.1051 | -106.2353
20.1088 | -106.2218
20.1125 | -106.2088
20.1162 | -106.195
20.1199 | -106.1806
20.1236 | -106.1667
20.1272 | -106.1527
20.1309 | -106.138
20.1346 | -106.1235
20.1383 | -106.109
20.142 | -106.0936
20.1457 | -106.0775
20.1494 | -106.0638
20.1531 | -106.0514
20.1568 | -106.0369
20.1604 | -106.0228
20.1641 | -106.0098
20.1678 | -105.9962
20.1715 | -105.983
20.1752 | -105.9703
20.1789 | -105.9575
20.1826 | -105.9449
20.1863 | -105.9345
20.1899 | -105.9238
20.1936 | -105.9143
20.1973 | -105.9065
20.201 | -105.8977
20.2047 | -105.8871
20.2084 | -105.8775
20.2121 | -105.8705
20.2158 | -105.8662
20.2195 | -105.8679
20.2231 | -105.8627
20.2268 | -105.8562
20.2305 | -105.8513
20.2342 | -105.8471
20.2379 | -105.8451
20.2416 | -105.845
20.2453 | -105.8464
20.249 | -105.8486
20.2527 | -105.8503
20.2563 | -105.8522
20.26 | -105.8536
20.2637 | -105.8547
20.2674 | -105.8549
20.2711 | -105.8585
20.2748 | -105.8632
20.2785 | -105.868
20.2822 | -105.8699
20.2858 | -105.8715
20.2895 | -105.8762
20.2932 | -105.8805
20.2969 | -105.8839
20.3006 | -105.8872
20.3043 | -105.8887
20.308 | -105.8928
20.3117 | -105.8973
20.3154 | -105.902
20.319 | -105.9055
20.3227 | -105.909
20.3264 | -105.9132
20.3301 | -105.9182
20.3338 | -105.9252
20.3375 | -105.9343
20.3412 | -105.9411
20.3449 | -105.947
20.3485 | -105.9525
20.3522 | -105.9575
20.3559 | -105.9641
20.3596 | -105.9718
20.3633 | -105.9806
20.367 | -105.9861
20.3707 | -105.9913
20.3744 | -105.9976
20.3781 | -106.0038
20.3817 | -106.0092
20.3854 | -106.0108
20.3891 | -106.0157
20.3928 | -106.0224
20.3965 | -106.0284
20.4002 | -106.0318
20.4039 | -106.0356
20.4076 | -106.0398
20.4112 | -106.041
20.4149 | -106.0423
20.4186 | -106.0438
20.4223 | -106.0459
20.426 | -106.0473
20.4297 | -106.0468
20.4334 | -106.0466
20.4371 | -106.0447
20.4408 | -106.0414
20.4444 | -106.0407
20.4481 | -106.0395
20.4518 | -106.0364
20.4555 | -106.0354
20.4592 | -106.0341
20.4629 | -106.03
20.4666 | -106.0266
20.4703 | -106.0241
20.474 | -106.0215
20.4776 | -106.0154
20.4813 | -106.0084
20.485 | -106.0028
20.4887 | -105.9996
20.4924 | -105.9969
20.4961 | -105.989
20.4998 | -105.98
20.5035 | -105.9731
20.5071 | -105.9676
20.5108 | -105.9619
20.5145 | -105.9567
20.5182 | -105.9509
20.5219 | -105.9461
20.5256 | -105.9411
20.5293 | -105.9355
20.533 | -105.9306
20.5367 | -105.9245
20.5403 | -105.9189
20.544 | -105.914
20.5477 | -105.9101
20.5514 | -105.9067
20.5551 | -105.9039
20.5588 | -105.9024
20.5625 | -105.9006
20.5662 | -105.8958
20.5698 | -105.8925
20.5735 | -105.8915
20.5772 | -105.8909
20.5809 | -105.8883
20.5846 | -105.886
20.5883 | -105.8847
20.592 | -105.8823
20.5957 | -105.8807
20.5994 | -105.8791
20.603 | -105.8802
20.6067 | -105.8813
20.6104 | -105.8816
20.6141 | -105.8819
20.6178 | -105.8822
20.6215 | -105.8817
20.6252 | -105.8803
20.6289 | -105.8798
20.6325 | -105.8802
20.6362 | -105.8804
20.6399 | -105.8797
20.6436 | -105.878
20.6473 | -105.875
20.651 | -105.8706
20.6547 | -105.8659
20.6584 | -105.8645
20.6621 | -105.8638
20.6657 | -105.8618
20.6694 | -105.8579
20.6731 | -105.8551
20.6768 | -105.8526
20.6805 | -105.8483
20.6842 | -105.8434
20.6879 | -105.8401
20.6916 | -105.8369
20.6953 | -105.8336
20.6989 | -105.8319
20.7026 | -105.8305
20.7063 | -105.828
20.71 | -105.8246
20.7137 | -105.8213
20.7174 | -105.8196
20.7211 | -105.8203
20.7248 | -105.8203
20.7284 | -105.8195
20.7321 | -105.8177
20.7358 | -105.814
20.7395 | -105.8132
20.7432 | -105.8136
20.7469 | -105.8137
20.7506 | -105.8189
20.7543 | -105.82
20.758 | -105.8222
20.7616 | -105.8259
20.7653 | -105.8292
20.769 | -105.8306
20.7727 | -105.8353
20.7764 | -105.8407
20.7801 | -105.843
20.7838 | -105.8434
20.7875 | -105.845
20.7911 | -105.8498
20.7948 | -105.8551
20.7985 | -105.8563
20.8022 | -105.8553
20.8059 | -105.8558
20.8096 | -105.8573
20.8133 | -105.8601
20.817 | -105.8618
20.8207 | -105.8615
20.8243 | -105.86
20.828 | -105.859
20.8317 | -105.8597
20.8354 | -105.859
20.8391 | -105.8603
20.8428 | -105.8618
20.8465 | -105.8615
20.8502 | -105.8608
20.8539 | -105.8606
20.8575 | -105.8597
20.8612 | -105.8615
20.8649 | -105.8596
20.8686 | -105.8579
20.8723 | -105.859
20.876 | -105.8597
20.8797 | -105.8611
20.8834 | -105.8612
20.887 | -105.8628
20.8907 | -105.8626
20.8944 | -105.862
20.8981 | -105.8619
20.9018 | -105.8638
20.9055 | -105.8676
20.9092 | -105.872
20.9129 | -105.8754
20.9166 | -105.8796
20.9202 | -105.8846
20.9239 | -105.8865
20.9276 | -105.8874
20.9313 | -105.8897
20.935 | -105.8973
20.9387 | -105.9051
20.9424 | -105.9116
20.9461 | -105.9199
20.9497 | -105.9278
20.9534 | -105.933
20.9571 | -105.9363
20.9608 | -105.9396
20.9645 | -105.9435
20.9682 | -105.9491
20.9719 | -105.9557
20.9756 | -105.9589
20.9793 | -105.9625
20.9829 | -105.9654
20.9866 | -105.969
20.9903 | -105.9715
20.994 | -105.9725
20.9977 | -105.9734
21.0014 | -105.9743
21.0051 | -105.9746
21.0088 | -105.9716
21.0124 | -105.9682
21.0161 | -105.9653
21.0198 | -105.9616
21.0235 | -105.9592
21.0272 | -105.956
21.0309 | -105.9522
21.0346 | -105.9452
21.0383 | -105.9356
21.042 | -105.9261
21.0456 | -105.9162
21.0493 | -105.9089
21.053 | -105.9004
21.0567 | -105.8918
21.0604 | -105.8832
21.0641 | -105.8722
21.0678 | -105.8632
21.0715 | -105.8548
21.0752 | -105.8465
21.0788 | -105.8395
21.0825 | -105.833
21.0862 | -105.8263
21.0899 | -105.8205
21.0936 | -105.8143
21.0973 | -105.8102
21.101 | -105.8078
21.1047 | -105.8038
21.1083 | -105.7994
21.112 | -105.7957
21.1157 | -105.7933
21.1194 | -105.7949
21.1231 | -105.7972
21.1268 | -105.7987
21.1305 | -105.8009
21.1342 | -105.8043
21.1379 | -105.8068
21.1415 | -105.8093
21.1452 | -105.8145
21.1489 | -105.8198
21.1526 | -105.8263
21.1563 | -105.8332
21.16 | -105.839
21.1637 | -105.8485
21.1674 | -105.8575
21.171 | -105.8658
21.1747 | -105.8752
21.1784 | -105.8814
21.1821 | -105.887
21.1858 | -105.8948
21.1895 | -105.9066
21.1932 | -105.9153
21.1969 | -105.9209
21.2006 | -105.9273
21.2042 | -105.934
21.2079 | -105.9385
21.2116 | -105.9445
21.2153 | -105.9497
21.219 | -105.9543
21.2227 | -105.9617
21.2264 | -105.9684
21.2301 | -105.9727
21.2337 | -105.9717
21.2374 | -105.968
21.2411 | -105.9657
21.2448 | -105.961
21.2485 | -105.9572
21.2522 | -105.9552
21.2559 | -105.9541
21.2596 | -105.9504
21.2633 | -105.9457
21.2669 | -105.9395
21.2706 | -105.9329
21.2743 | -105.9263
21.278 | -105.9194
21.2817 | -105.9115
21.2854 | -105.9052
21.2891 | -105.8962
21.2928 | -105.8867
21.2965 | -105.8781
21.3001 | -105.8702
21.3038 | -105.8627
21.3075 | -105.8556
21.3112 | -105.8494
21.3149 | -105.8419
21.3186 | -105.8331
21.3223 | -105.8253
21.326 | -105.8186
21.3296 | -105.8145
21.3333 | -105.8111
21.337 | -105.8076
21.3407 | -105.8041
21.3444 | -105.8002
21.3481 | -105.7979
21.3518 | -105.7958
21.3555 | -105.7955
21.3592 | -105.7955
21.3628 | -105.7932
21.3665 | -105.7943
21.3702 | -105.7961
21.3739 | -105.7972
21.3776 | -105.7993
21.3813 | -105.8036
21.385 | -105.8092
21.3887 | -105.8154
21.3923 | -105.8218
21.396 | -105.8276
21.3997 | -105.8329
21.4034 | -105.8393
21.4071 | -105.8489
21.4108 | -105.8598
21.4145 | -105.8705
21.4182 | -105.8816
21.4219 | -105.89
21.4255 | -105.8988
21.4292 | -105.9085
21.4329 | -105.9184
21.4366 | -105.9287
21.4403 | -105.938
21.444 | -105.9489
21.4477 | -105.9599
21.4514 | -105.9713
21.455 | -105.9807
21.4587 | -105.9907
21.4624 | -106.0021
21.4661 | -106.0128
21.4698 | -106.0213
21.4735 | -106.0277
21.4772 | -106.0335
21.4809 | -106.0397
21.4846 | -106.0479
21.4882 | -106.0546
21.4919 | -106.0589
21.4956 | -106.0625
21.4993 | -106.0661
21.503 | -106.0686
21.5067 | -106.069
21.5104 | -106.0681
21.5141 | -106.064
21.5177 | -106.0596
21.5214 | -106.0553
21.5251 | -106.0481
21.5288 | -106.0412
21.5325 | -106.0355
21.5362 | -106.0293
21.5399 | -106.0226
21.5436 | -106.0126
21.5473 | -106.0007
21.5509 | -105.9867
21.5546 | -105.9693
21.5583 | -105.9535
21.562 | -105.9378
21.5657 | -105.9216
21.5694 | -105.9057
21.5731 | -105.8904
21.5768 | -105.8763
21.5805 | -105.8624
21.5841 | -105.8476
21.5878 | -105.8315
21.5915 | -105.8159
21.5952 | -105.8025
21.5989 | -105.7902
21.6026 | -105.777
21.6063 | -105.7639
21.61 | -105.7504
21.6136 | -105.7368
21.6173 | -105.7255
21.621 | -105.7153
21.6247 | -105.7054
21.6284 | -105.6962
21.6321 | -105.6867
21.6358 | -105.6758
21.6395 | -105.665
21.6432 | -105.6563
21.6468 | -105.6502
21.6505 | -105.644
21.6542 | -105.6367
21.6579 | -105.629
21.6616 | -105.6229
21.6653 | -105.6187
21.669 | -105.6163
21.6727 | -105.6141
21.6763 | -105.6131
21.68 | -105.6116
21.6837 | -105.6089
21.6874 | -105.6047
21.6911 | -105.6029
21.6948 | -105.6033
21.6985 | -105.6024
21.7022 | -105.6036
21.7059 | -105.6038
21.7095 | -105.6034
21.7132 | -105.604
21.7169 | -105.6054
21.7206 | -105.6071
21.7243 | -105.6075
21.728 | -105.6088
21.7317 | -105.6089
21.7354 | -105.6107
21.739 | -105.6108
21.7427 | -105.6104
21.7464 | -105.6125
21.7501 | -105.6144
21.7538 | -105.6146
21.7575 | -105.6122
21.7612 | -105.612
21.7649 | -105.6122
21.7686 | -105.6095
21.7722 | -105.6091
21.7759 | -105.6099
21.7796 | -105.6096
21.7833 | -105.6094
21.787 | -105.6093
21.7907 | -105.607
21.7944 | -105.6026
21.7981 | -105.599
21.8018 | -105.5976
21.8054 | -105.5963
21.8091 | -105.5923
21.8128 | -105.5895
21.8165 | -105.5836
21.8202 | -105.5815
21.8239 | -105.5798
21.8276 | -105.5771
21.8313 | -105.5744
21.8349 | -105.5724
21.8386 | -105.5691
21.8423 | -105.5663
21.846 | -105.5643
21.8497 | -105.5628
21.8534 | -105.5616
21.8571 | -105.56
21.8608 | -105.5568
21.8645 | -105.557
21.8681 | -105.5551
21.8718 | -105.5542
21.8755 | -105.5525
21.8792 | -105.5489
21.8829 | -105.5451
21.8866 | -105.543
21.8903 | -105.5395
21.894 | -105.5372
21.8976 | -105.5363
21.9013 | -105.5355
21.905 | -105.533
21.9087 | -105.5312
21.9124 | -105.5298
21.9161 | -105.5291
21.9198 | -105.5295
21.9235 | -105.5318
21.9272 | -105.5323
21.9308 | -105.534
21.9345 | -105.5342
21.9382 | -105.5335
21.9419 | -105.5336
21.9456 | -105.5349
21.9493 | -105.5355
21.953 | -105.5335
21.9567 | -105.5317
21.9603 | -105.5295
21.964 | -105.5302
21.9677 | -105.5307
21.9714 | -105.5308
21.9751 | -105.5295
21.9788 | -105.53
21.9825 | -105.5279
21.9862 | -105.5246
21.9899 | -105.5248
21.9935 | -105.5258
21.9972 | -105.5242
22.0009 | -105.5199
22.0046 | -105.5155
22.0083 | -105.5133
22.012 | -105.5108
22.0157 | -105.5083
22.0194 | -105.505
22.0231 | -105.5007
22.0267 | -105.4951
22.0304 | -105.4909
22.0341 | -105.4864
22.0378 | -105.4821
22.0415 | -105.4796
22.0452 | -105.4755
22.0489 | -105.472
22.0526 | -105.4683
22.0562 | -105.4647
22.0599 | -105.4593
22.0636 | -105.4561
22.0673 | -105.4554
22.071 | -105.4529
22.0747 | -105.4493
22.0784 | -105.4467
22.0821 | -105.4409
22.0858 | -105.4375
22.0894 | -105.4355
22.0931 | -105.4324
22.0968 | -105.4301
22.1005 | -105.4292
22.1042 | -105.431
22.1079 | -105.4331
22.1116 | -105.4352
22.1153 | -105.4356
22.1189 | -105.4374
22.1226 | -105.4396
22.1263 | -105.4439
22.13 | -105.4467
22.1337 | -105.4496
22.1374 | -105.453
22.1411 | -105.4557
22.1448 | -105.4585
22.1485 | -105.4613
22.1521 | -105.4659
22.1558 | -105.4686
22.1595 | -105.4725
22.1632 | -105.4789
22.1669 | -105.4839
22.1706 | -105.4896
22.1743 | -105.4973
22.178 | -105.5027
22.1817 | -105.5077
22.1853 | -105.5103
22.189 | -105.5139
22.1927 | -105.5177
22.1964 | -105.5202
22.2001 | -105.5234
22.2038 | -105.5263
22.2075 | -105.5265
22.2112 | -105.5281
22.2148 | -105.5298
22.2185 | -105.5325
22.2222 | -105.5295
22.2259 | -105.5278
22.2296 | -105.5262
22.2333 | -105.5234
22.237 | -105.5202
22.2407 | -105.5174
22.2444 | -105.5139
22.248 | -105.5131
22.2517 | -105.512
22.2554 | -105.5104
22.2591 | -105.5084
22.2628 | -105.5067
22.2665 | -105.5059
22.2702 | -105.5039
22.2739 | -105.5019
22.2775 | -105.4965
22.2812 | -105.4916
22.2849 | -105.4866
22.2886 | -105.4825
22.2923 | -105.4767
22.296 | -105.4725
22.2997 | -105.4683
22.3034 | -105.4635
22.3071 | -105.4603
22.3107 | -105.4596
22.3144 | -105.4573
22.3181 | -105.4528
22.3218 | -105.4469
22.3255 | -105.4422
22.3292 | -105.4392
22.3329 | -105.4372
22.3366 | -105.4335
22.3402 | -105.4286
22.3439 | -105.4227
22.3476 | -105.4187
22.3513 | -105.4161
22.355 | -105.413
22.3587 | -105.4089
22.3624 | -105.4044
22.3661 | -105.3995
22.3698 | -105.395
22.3734 | -105.3896
22.3771 | -105.3847
22.3808 | -105.3802
22.3845 | -105.3756
22.3882 | -105.3702
22.3919 | -105.3644
22.3956 | -105.3607
22.3993 | -105.3569
22.403 | -105.3506
22.4066 | -105.3436
22.4103 | -105.3364
22.414 | -105.3298
22.4177 | -105.323
22.4214 | -105.3136
22.4251 | -105.3031
22.4288 | -105.2944
22.4325 | -105.2873
22.4361 | -105.2802
22.4398 | -105.273
22.4435 | -105.2652
22.4472 | -105.2583
22.4509 | -105.2519
22.4546 | -105.244
22.4583 | -105.2356
22.462 | -105.2275
22.4657 | -105.2201
22.4693 | -105.2127
22.473 | -105.2066
22.4767 | -105.1996
22.4804 | -105.1916
22.4841 | -105.1849
22.4878 | -105.1769
22.4915 | -105.1684
22.4952 | -105.1621
22.4988 | -105.1571
22.5025 | -105.1524
22.5062 | -105.1463
22.5099 | -105.1417
22.5136 | -105.1395
22.5173 | -105.1364
22.521 | -105.1339
22.5247 | -105.1312
22.5284 | -105.1293
22.532 | -105.126
22.5357 | -105.1221
22.5394 | -105.1176
22.5431 | -105.111
22.5468 | -105.1012
22.5505 | -105.0832
22.5542 | -105.0636
22.5579 | -105.0458
22.5615 | -105.0234
22.5652 | -104.9997
22.5689 | -104.9737
22.5726 | -104.9457
22.5763 | -104.9137
22.58 | -104.8824
22.5837 | -104.8509
22.5874 | -104.8151
22.5911 | -104.7795
22.5947 | -104.7409
22.5984 | -104.6999
22.6021 | -104.6577
22.6058 | -104.6119
22.6095 | -104.5626
22.6132 | -104.5107
22.6169 | -104.4572
22.6206 | -104.4014
22.6243 | -104.3433
22.6279 | -104.2836
22.6316 | -104.2243
22.6353 | -104.1629
22.639 | -104.0999
22.6427 | -104.0346
22.6464 | -103.9701
22.6501 | -103.9056
22.6538 | -103.8425
22.6574 | -103.7799
22.6611 | -103.7167
22.6648 | -103.6531
22.6685 | -103.591
22.6722 | -103.5308
22.6759 | -103.4722
22.6796 | -103.4166
22.6833 | -103.364
22.687 | -103.3137
22.6906 | -103.2668
22.6943 | -103.2233
22.698 | -103.1832
22.7017 | -103.1462
22.7054 | -103.1129
22.7091 | -103.0837
22.7128 | -103.059
22.7165 | -103.0377
22.7201 | -103.021
22.7238 | -103.008
22.7275 | -103
22.7312 | -102.9975
22.7349 | -102.9983
22.7386 | -103.0026
22.7423 | -103.0119
22.746 | -103.0256
22.7497 | -103.0428
22.7533 | -103.0641
22.757 | -103.0885
22.7607 | -103.1165
22.7644 | -103.1486
22.7681 | -103.1847
22.7718 | -103.2235
22.7755 | -103.2654
22.7792 | -103.3115
22.7828 | -103.3594
22.7865 | -103.4086
22.7902 | -103.46
22.7939 | -103.5136
22.7976 | -103.5697
22.8013 | -103.6272
22.805 | -103.6846
22.8087 | -103.7435
22.8124 | -103.8036
22.816 | -103.8629
22.8197 | -103.9203
22.8234 | -103.9798
22.8271 | -104.0408
22.8308 | -104.1
22.8345 | -104.1559
22.8382 | -104.2106
22.8419 | -104.2638
22.8456 | -104.3171
22.8492 | -104.3674
22.8529 | -104.4168
22.8566 | -104.4641
22.8603 | -104.512
22.864 | -104.558
22.8677 | -104.6025
22.8714 | -104.6498
22.8751 | -104.6938
22.8787 | -104.7096
22.8824 | -104.7457
22.8861 | -104.7806
22.8898 | -104.8115
22.8935 | -104.8428
22.8972 | -104.8705
22.9009 | -104.8975
22.9046 | -104.9194
22.9083 | -104.9362
22.9119 | -104.9515
22.9156 | -104.9633
22.9193 | -104.9755
22.923 | -104.9872
22.9267 | -104.9955
22.9304 | -105.0023
22.9341 | -105.0115
22.9378 | -105.0209
22.9414 | -105.028
22.9451 | -105.0336
22.9488 | -105.0379
22.9525 | -105.0394
22.9562 | -105.0419
22.9599 | -105.0454
22.9636 | -105.0498
22.9673 | -105.0531
22.971 | -105.0544
22.9746 | -105.0547
22.9783 | -105.0518
22.982 | -105.0509
22.9857 | -105.0524
22.9894 | -105.052
22.9931 | -105.0497
22.9968 | -105.0464
23.0005 | -105.0432
23.0041 | -105.0383
23.0078 | -105.0338
23.0115 | -105.0284
23.0152 | -105.0229
23.0189 | -105.0166
23.0226 | -105.0101
23.0263 | -105.0024
23.03 | -104.9942
23.0337 | -104.9871
23.0373 | -104.9789
23.041 | -104.9699
23.0447 | -104.9604
23.0484 | -104.9524
23.0521 | -104.9436
23.0558 | -104.9343
23.0595 | -104.9244
23.0632 | -104.9152
23.0669 | -104.9066
23.0705 | -104.8988
23.0742 | -104.8916
23.0779 | -104.8822
23.0816 | -104.874
23.0853 | -104.868
23.089 | -104.8604
23.0927 | -104.8524
23.0964 | -104.8446
23.1 | -104.8388
23.1037 | -104.8327
23.1074 | -104.8264
23.1111 | -104.8202
23.1148 | -104.8142
23.1185 | -104.8099
23.1222 | -104.806
23.1259 | -104.802
23.1296 | -104.7991
23.1332 | -104.7949
23.1369 | -104.7918
23.1406 | -104.7887
23.1443 | -104.7879
23.148 | -104.7858
23.1517 | -104.7829
23.1554 | -104.7793
23.1591 | -104.7775
23.1627 | -104.775
23.1664 | -104.7723
23.1701 | -104.7708
23.1738 | -104.7707
23.1775 | -104.7704
23.1812 | -104.7704
23.1849 | -104.7688
23.1886 | -104.7657
23.1923 | -104.7647
23.1959 | -104.7653
23.1996 | -104.7646
23.2033 | -104.7648
23.207 | -104.7627
23.2107 | -104.7609
23.2144 | -104.7597
23.2181 | -104.7582
23.2218 | -104.7556
23.2254 | -104.754
23.2291 | -104.7524
23.2328 | -104.7508
23.2365 | -104.7508
23.2402 | -104.7512
23.2439 | -104.7503
23.2476 | -104.749
23.2513 | -104.7475
23.255 | -104.7461
23.2586 | -104.7459
23.2623 | -104.7432
23.266 | -104.7393
23.2697 | -104.7345
23.2734 | -104.7305
23.2771 | -104.7282
23.2808 | -104.7273
23.2845 | -104.7228
23.2882 | -104.7145
23.2918 | -104.7064
23.2955 | -104.701
23.2992 | -104.6969
23.3029 | -104.6922
23.3066 | -104.6884
23.3103 | -104.6823
23.314 | -104.6744
23.3177 | -104.6682
23.3213 | -104.663
23.325 | -104.6573
23.3287 | -104.6525
23.3324 | -104.6478
23.3361 | -104.6416
23.3398 | -104.637
23.3435 | -104.6309
23.3472 | -104.6245
23.3509 | -104.6193
23.3545 | -104.6151
23.3582 | -104.6121
23.3619 | -104.6077
23.3656 | -104.6023
23.3693 | -104.5995
23.373 | -104.596
23.3767 | -104.5931
23.3804 | -104.5906
23.384 | -104.5883
23.3877 | -104.5867
23.3914 | -104.5857
23.3951 | -104.5836
23.3988 | -104.5815
23.4025 | -104.5798
23.4062 | -104.5803
23.4099 | -104.5832
23.4136 | -104.584
23.4172 | -104.5841
23.4209 | -104.5859
23.4246 | -104.5889
23.4283 | -104.5921
23.432 | -104.5947
23.4357 | -104.5969
23.4394 | -104.5994
23.4431 | -104.6032
23.4468 | -104.6063
23.4504 | -104.608
23.4541 | -104.6112
23.4578 | -104.6155
23.4615 | -104.6209
23.4652 | -104.6277
23.4689 | -104.6338
23.4726 | -104.6395
23.4763 | -104.6459
23.4799 | -104.6521
23.4836 | -104.6587
23.4873 | -104.6667
23.491 | -104.6757
23.4947 | -104.6869
23.4984 | -104.6984
23.5021 | -104.7088
23.5058 | -104.7191
23.5095 | -104.7307
23.5131 | -104.7416
23.5168 | -104.7553
23.5205 | -104.77
23.5242 | -104.7824
23.5279 | -104.7931
23.5316 | -104.8058
23.5353 | -104.8191
23.539 | -104.8347
23.5426 | -104.8494
23.5463 | -104.8627
23.55 | -104.8764
23.5537 | -104.89
23.5574 | -104.9035
23.5611 | -104.9165
23.5648 | -104.9311
23.5685 | -104.9441
23.5722 | -104.9562
23.5758 | -104.9696
23.5795 | -104.9828
23.5832 | -104.9953
23.5869 | -105.0081
23.5906 | -105.0215
23.5943 | -105.0355
23.598 | -105.0501
23.6017 | -105.0655
23.6053 | -105.079
23.609 | -105.0901
23.6127 | -105.102
23.6164 | -105.114
23.6201 | -105.1283
23.6238 | -105.1405
23.6275 | -105.1506
23.6312 | -105.1592
23.6349 | -105.1678
23.6385 | -105.1767
23.6422 | -105.1869
23.6459 | -105.1976
23.6496 | -105.2063
23.6533 | -105.2134
23.657 | -105.2192
23.6607 | -105.2252
23.6644 | -105.2314
23.6681 | -105.2383
23.6717 | -105.2445
23.6754 | -105.2507
23.6791 | -105.2563
23.6828 | -105.2624
23.6865 | -105.2683
23.6902 | -105.2727
23.6939 | -105.2785
23.6976 | -105.2856
23.7012 | -105.2908
23.7049 | -105.2965
23.7086 | -105.2991
23.7123 | -105.3021
23.716 | -105.3059
23.7197 | -105.3107
23.7234 | -105.3159
23.7271 | -105.3219
23.7308 | -105.3259
23.7344 | -105.3299
23.7381 | -105.3352
23.7418 | -105.3415
23.7455 | -105.3474
23.7492 | -105.3545
23.7529 | -105.3608
23.7566 | -105.3683
23.7603 | -105.3755
23.7639 | -105.3839
23.7676 | -105.3918
23.7713 | -105.4022
23.775 | -105.4135
23.7787 | -105.4231
23.7824 | -105.4329
23.7861 | -105.4433
23.7898 | -105.4522
23.7935 | -105.4623
23.7971 | -105.4725
23.8008 | -105.4825
23.8045 | -105.4905
23.8082 | -105.5005
23.8119 | -105.5104
23.8156 | -105.5218
23.8193 | -105.5329
23.823 | -105.545
23.8266 | -105.5548
23.8303 | -105.5617
23.834 | -105.57
23.8377 | -105.5797
23.8414 | -105.5891
23.8451 | -105.5966
23.8488 | -105.606
23.8525 | -105.6139
23.8562 | -105.6208
23.8598 | -105.6282
23.8635 | -105.6352
23.8672 | -105.6405
23.8709 | -105.6433
23.8746 | -105.6448
23.8783 | -105.6483
23.882 | -105.6528
23.8857 | -105.6562
23.8894 | -105.6577
23.893 | -105.6584
23.8967 | -105.6582
23.9004 | -105.6551
23.9041 | -105.6521
23.9078 | -105.6499
23.9115 | -105.6463
23.9152 | -105.6398
23.9189 | -105.6321
23.9225 | -105.6253
23.9262 | -105.6174
23.9299 | -105.6071
23.9336 | -105.5963
23.9373 | -105.5863
23.941 | -105.5771
23.9447 | -105.562
23.9484 | -105.5526
23.9521 | -105.5418
23.9557 | -105.5311
23.9594 | -105.5207
23.9631 | -105.5097
23.9668 | -105.4956
23.9705 | -105.4806
23.9742 | -105.4674
23.9779 | -105.4555
23.9816 | -105.4444
23.9852 | -105.4334
23.9889 | -105.4235
23.9926 | -105.4139
23.9963 | -105.4051
24 | -105.3944
24.0037 | -105.3844
24.0074 | -105.3753
24.0111 | -105.368
24.0148 | -105.3598
24.0184 | -105.3528
24.0221 | -105.3469
24.0258 | -105.3412
24.0295 | -105.3352
24.0332 | -105.3305
24.0369 | -105.3269
24.0406 | -105.3219
24.0443 | -105.3188
24.0479 | -105.3168
24.0516 | -105.314
24.0553 | -105.3117
24.059 | -105.3101
24.0627 | -105.3103
24.0664 | -105.3115
24.0701 | -105.3148
24.0738 | -105.317
24.0775 | -105.3185
24.0811 | -105.3184
24.0848 | -105.3207
24.0885 | -105.3224
24.0922 | -105.3234
24.0959 | -105.3252
24.0996 | -105.3285
24.1033 | -105.3321
24.107 | -105.3366
24.1106 | -105.3415
24.1143 | -105.3458
24.118 | -105.3511
24.1217 | -105.3571
24.1254 | -105.3612
24.1291 | -105.3639
24.1328 | -105.3669
24.1365 | -105.3695
24.1402 | -105.3718
24.1438 | -105.3746
24.1475 | -105.3767
24.1512 | -105.3791
24.1549 | -105.3817
24.1586 | -105.3839
24.1623 | -105.3848
24.166 | -105.3858
24.1697 | -105.3881
24.1734 | -105.3894
24.177 | -105.3896
24.1807 | -105.3895
24.1844 | -105.3911
24.1881 | -105.3922
24.1918 | -105.3934
24.1955 | -105.3937
24.1992 | -105.3937
24.2029 | -105.3944
24.2065 | -105.3942
24.2102 | -105.3936
24.2139 | -105.3938
24.2176 | -105.3906
24.2213 | -105.3885
24.225 | -105.388
24.2287 | -105.3864
24.2324 | -105.3851
24.2361 | -105.3855
24.2397 | -105.3867
24.2434 | -105.3863
24.2471 | -105.3855
24.2508 | -105.3847
24.2545 | -105.3857
24.2582 | -105.3869
24.2619 | -105.3872
24.2656 | -105.3866
24.2692 | -105.3873
24.2729 | -105.3876
24.2766 | -105.3874
24.2803 | -105.3847
24.284 | -105.3829
24.2877 | -105.3807
24.2914 | -105.3765
24.2951 | -105.3685
24.2988 | -105.3559
24.3024 | -105.3384
24.3061 | -105.3264
24.3098 | -105.317
24.3135 | -105.307
24.3172 | -105.2939
24.3209 | -105.2811
24.3246 | -105.2673
24.3283 | -105.2529
24.3319 | -105.2391
24.3356 | -105.225
24.3393 | -105.2092
24.343 | -105.1909
24.3467 | -105.172
24.3504 | -105.1522
24.3541 | -105.1295
24.3578 | -105.1073
24.3615 | -105.0847
24.3651 | -105.0596
24.3688 | -105.034
24.3725 | -105.0096
24.3762 | -104.9835
24.3799 | -104.955
24.3836 | -104.9253
24.3873 | -104.8944
24.391 | -104.8655
24.3947 | -104.8367
24.3983 | -104.8064
24.402 | -104.7758
24.4057 | -104.7445
24.4094 | -104.7141
24.4131 | -104.6827
24.4168 | -104.6531
24.4205 | -104.6246
24.4242 | -104.5977
24.4278 | -104.5706
24.4315 | -104.5455
24.4352 | -104.5236
24.4389 | -104.5004
24.4426 | -104.477
24.4463 | -104.4571
24.45 | -104.4388
24.4537 | -104.4212
24.4574 | -104.4067
24.461 | -104.3957
24.4647 | -104.3876
24.4684 | -104.3819
24.4721 | -104.3783
24.4758 | -104.3933
24.4795 | -104.3924
24.4832 | -104.3934
24.4869 | -104.3956
24.4905 | -104.3984
24.4942 | -104.4028
24.4979 | -104.4073
24.5016 | -104.4138
24.5053 | -104.4222
24.509 | -104.4336
24.5127 | -104.4475
24.5164 | -104.4615
24.5201 | -104.476
24.5237 | -104.4908
24.5274 | -104.5059
24.5311 | -104.5207
24.5348 | -104.5373
24.5385 | -104.5542
24.5422 | -104.5715
24.5459 | -104.5906
24.5496 | -104.609
24.5532 | -104.6277
24.5569 | -104.647
24.5606 | -104.6652
24.5643 | -104.6821
24.568 | -104.6953
24.5717 | -104.6994
24.5754 | -104.6999
24.5791 | -104.6971
24.5828 | -104.6881
24.5864 | -104.6759
24.5901 | -104.649
24.5938 | -104.597
24.5975 | -104.5413
24.6012 | -104.4872
24.6049 | -104.4372
24.6086 | -104.3904
24.6123 | -104.341
24.616 | -104.287
24.6196 | -104.2267
24.6233 | -104.1632
24.627 | -104.1005
24.6307 | -104.0355
24.6344 | -103.9668
24.6381 | -103.8958
24.6418 | -103.8226
24.6455 | -103.7481
24.6491 | -103.6721
24.6528 | -103.588
24.6565 | -103.4993
24.6602 | -103.4078
24.6639 | -103.3147
24.6676 | -103.2195
24.6713 | -103.123
24.675 | -103.0258
24.6787 | -102.928
24.6823 | -102.8302
24.686 | -102.7329
24.6897 | -102.6371
24.6934 | -102.5435
24.6971 | -102.4522
24.7008 | -102.363
24.7045 | -102.2756
24.7082 | -102.1906
24.7118 | -102.1087
24.7155 | -102.0309
24.7192 | -101.9565
24.7229 | -101.8868
24.7266 | -101.8213
24.7303 | -101.7601
24.734 | -101.7034
24.7377 | -101.6511
24.7414 | -101.6045
24.745 | -101.5633
24.7487 | -101.5274
24.7524 | -101.4962
24.7561 | -101.4704
24.7598 | -101.4515
24.7635 | -101.4384
24.7672 | -101.4306
24.7709 | -101.4284
24.7746 | -101.4319
24.7782 | -101.4411
24.7819 | -101.4558
24.7856 | -101.4774
24.7893 | -101.5047
24.793 | -101.5362
24.7967 | -101.5752
24.8004 | -101.62
24.8041 | -101.6693
24.8077 | -101.7236
24.8114 | -101.7822
24.8151 | -101.8455
24.8188 | -101.914
24.8225 | -101.9874
24.8262 | -102.0644
24.8299 | -102.1453
24.8336 | -102.2305
24.8373 | -102.3194
24.8409 | -102.4113
24.8446 | -102.506
24.8483 | -102.6041
24.852 | -102.7031
24.8557 | -102.8021
24.8594 | -102.9019
24.8631 | -103.0022
24.8668 | -103.1036
24.8704 | -103.2058
24.8741 | -103.3076
24.8778 | -103.4087
24.8815 | -103.5096
24.8852 | -103.6077
24.8889 | -103.704
24.8926 | -103.7973
24.8963 | -103.8873
24.9 | -103.9757
24.9036 | -104.0608
24.9073 | -104.143
24.911 | -104.2229
24.9147 | -104.3007
24.9184 | -104.3763
24.9221 | -104.4547
24.9258 | -104.5305
24.9295 | -104.6012
24.9331 | -104.6679
24.9368 | -104.7303
24.9405 | -104.8004
24.9442 | -104.8842
24.9479 | -104.9558
24.9516 | -105.0113
24.9553 | -105.052
24.959 | -105.0804
24.9627 | -105.104
24.9663 | -105.1248
24.97 | -105.1442
24.9737 | -105.1591
24.9774 | -105.1665
24.9811 | -105.1711
24.9848 | -105.1744
24.9885 | -105.1729
24.9922 | -105.172
24.9959 | -105.1702
24.9995 | -105.1671
25.0032 | -105.1642
25.0069 | -105.1576
25.0106 | -105.1564
25.0143 | -105.1548
25.018 | -105.1532
25.0217 | -105.1493
25.0254 | -105.1445
25.029 | -105.1397
25.0327 | -105.1354
25.0364 | -105.1318
25.0401 | -105.1287
25.0438 | -105.1263
25.0475 | -105.1177
25.0512 | -105.1089
25.0549 | -105.1024
25.0586 | -105.0962
25.0622 | -105.0904
25.0659 | -105.0851
25.0696 | -105.0802
25.0733 | -105.0756
25.077 | -105.0699
25.0807 | -105.0655
25.0844 | -105.0628
25.0881 | -105.0599
25.0917 | -105.0562
25.0954 | -105.0514
25.0991 | -105.0479
25.1028 | -105.0452
25.1065 | -105.0441
25.1102 | -105.044
25.1139 | -105.0433
25.1176 | -105.042
25.1213 | -105.0419
25.1249 | -105.0432
25.1286 | -105.0454
25.1323 | -105.048
25.136 | -105.0503
25.1397 | -105.0526
25.1434 | -105.0558
25.1471 | -105.0616
25.1508 | -105.065
25.1544 | -105.0672
25.1581 | -105.071
25.1618 | -105.0743
25.1655 | -105.0811
25.1692 | -105.0898
25.1729 | -105.0974
25.1766 | -105.105
25.1803 | -105.113
25.184 | -105.1219
25.1876 | -105.131
25.1913 | -105.1403
25.195 | -105.1504
25.1987 | -105.1597
25.2024 | -105.1681
25.2061 | -105.1779
25.2098 | -105.19
25.2135 | -105.2015
25.2172 | -105.2121
25.2208 | -105.2208
25.2245 | -105.2277
25.2282 | -105.2375
25.2319 | -105.2467
25.2356 | -105.2535
25.2393 | -105.2611
25.243 | -105.2691
25.2467 | -105.2763
25.2503 | -105.283
25.254 | -105.2896
25.2577 | -105.2955
25.2614 | -105.2995
25.2651 | -105.3032
25.2688 | -105.3076
25.2725 | -105.3097
25.2762 | -105.309
25.2799 | -105.309
25.2835 | -105.3085
25.2872 | -105.3093
25.2909 | -105.3103
25.2946 | -105.3107
25.2983 | -105.3086
25.302 | -105.3055
25.3057 | -105.3022
25.3094 | -105.2963
25.313 | -105.2908
25.3167 | -105.2855
25.3204 | -105.2797
25.3241 | -105.2736
25.3278 | -105.2665
25.3315 | -105.2594
25.3352 | -105.2519
25.3389 | -105.2451
25.3426 | -105.239
25.3462 | -105.2313
25.3499 | -105.222
25.3536 | -105.2114
25.3573 | -105.2028
25.361 | -105.1956
25.3647 | -105.1877
25.3684 | -105.1777
25.3721 | -105.1681
25.3757 | -105.1588
25.3794 | -105.1506
25.3831 | -105.1452
25.3868 | -105.1394
25.3905 | -105.134
25.3942 | -105.1288
25.3979 | -105.1248
25.4016 | -105.1212
25.4053 | -105.1209
25.4089 | -105.1204
25.4126 | -105.1192
25.4163 | -105.1181
25.42 | -105.1185
25.4237 | -105.1186
25.4274 | -105.1169
25.4311 | -105.116
25.4348 | -105.1174
25.4385 | -105.1204
25.4421 | -105.1244
25.4458 | -105.128
25.4495 | -105.1309
25.4532 | -105.1338
25.4569 | -105.1373
25.4606 | -105.1418
25.4643 | -105.145
25.468 | -105.1481
25.4716 | -105.1526
25.4753 | -105.1589
25.479 | -105.1651
25.4827 | -105.1715
25.4864 | -105.1778
25.4901 | -105.1847
25.4938 | -105.1933
25.4975 | -105.2
25.5012 | -105.2069
25.5048 | -105.2145
25.5085 | -105.2205
25.5122 | -105.228
25.5159 | -105.2364
25.5196 | -105.2442
25.5233 | -105.2491
25.527 | -105.254
25.5307 | -105.2596
25.5343 | -105.2649
25.538 | -105.2723
25.5417 | -105.2775
25.5454 | -105.2831
25.5491 | -105.2866
25.5528 | -105.2915
25.5565 | -105.2965
25.5602 | -105.3016
25.5639 | -105.3059
25.5675 | -105.3108
25.5712 | -105.317
25.5749 | -105.3229
25.5786 | -105.3287
25.5823 | -105.3363
25.586 | -105.3424
25.5897 | -105.3456
25.5934 | -105.3503
25.597 | -105.3572
25.6007 | -105.3648
25.6044 | -105.3709
25.6081 | -105.3772
25.6118 | -105.3832
25.6155 | -105.3883
25.6192 | -105.394
25.6229 | -105.4002
25.6266 | -105.4051
25.6302 | -105.4083
25.6339 | -105.4123
25.6376 | -105.4187
25.6413 | -105.4256
25.645 | -105.4335
25.6487 | -105.4394
25.6524 | -105.444
25.6561 | -105.4477
25.6598 | -105.4529
25.6634 | -105.457
25.6671 | -105.461
25.6708 | -105.4663
25.6745 | -105.4683
25.6782 | -105.4704
25.6819 | -105.4733
25.6856 | -105.4758
25.6893 | -105.4801
25.6929 | -105.4842
25.6966 | -105.4879
25.7003 | -105.4895
25.704 | -105.4929
25.7077 | -105.4956
25.7114 | -105.498
25.7151 | -105.4993
25.7188 | -105.4985
25.7225 | -105.4984
25.7261 | -105.499
25.7298 | -105.499
25.7335 | -105.4999
25.7372 | -105.4999
25.7409 | -105.5
25.7446 | -105.4984
25.7483 | -105.4986
25.752 | -105.497
25.7556 | -105.4943
25.7593 | -105.4918
25.763 | -105.4897
25.7667 | -105.4886
25.7704 | -105.4872
25.7741 | -105.4862
25.7778 | -105.4871
25.7815 | -105.4865
25.7852 | -105.4865
25.7888 | -105.487
25.7925 | -105.4862
25.7962 | -105.4845
25.7999 | -105.4829
25.8036 | -105.4821
25.8073 | -105.483
25.811 | -105.484
25.8147 | -105.4861
25.8183 | -105.4877
25.822 | -105.4883
25.8257 | -105.4914
25.8294 | -105.4935
25.8331 | -105.4944
25.8368 | -105.4956
25.8405 | -105.4978
25.8442 | -105.4986
25.8479 | -105.4997
25.8515 | -105.5025
25.8552 | -105.5032
25.8589 | -105.5053
25.8626 | -105.5088
25.8663 | -105.5119
25.87 | -105.5157
25.8737 | -105.5199
25.8774 | -105.5229
25.8811 | -105.5262
25.8847 | -105.53
25.8884 | -105.5332
25.8921 | -105.537
25.8958 | -105.5412
25.8995 | -105.5442
25.9032 | -105.5464
25.9069 | -105.5503
25.9106 | -105.5538
25.9142 | -105.5585
25.9179 | -105.5607
25.9216 | -105.5625
25.9253 | -105.5652
25.929 | -105.5694
25.9327 | -105.5713
25.9364 | -105.5724
25.9401 | -105.5742
25.9438 | -105.577
25.9474 | -105.5795
25.9511 | -105.5832
25.9548 | -105.5875
25.9585 | -105.5909
25.9622 | -105.5932
25.9659 | -105.594
25.9696 | -105.5953
25.9733 | -105.5978
25.9769 | -105.5997
25.9806 | -105.6021
25.9843 | -105.6048
25.988 | -105.607
25.9917 | -105.6091
25.9954 | -105.6122
25.9991 | -105.6161
26.0028 | -105.6187
26.0065 | -105.6187
26.0101 | -105.6188
26.0138 | -105.6206
26.0175 | -105.6232
26.0212 | -105.6252
26.0249 | -105.6262
26.0286 | -105.6291
26.0323 | -105.6318
26.036 | -105.6348
26.0397 | -105.6394
26.0433 | -105.6423
26.047 | -105.6436
26.0507 | -105.6462
26.0544 | -105.6486
26.0581 | -105.651
26.0618 | -105.6543
26.0655 | -105.6582
26.0692 | -105.6605
26.0728 | -105.6664
26.0765 | -105.672
26.0802 | -105.6782
26.0839 | -105.6841
26.0876 | -105.6885
26.0913 | -105.6899
26.095 | -105.6916
26.0987 | -105.6967
26.1024 | -105.7006
26.106 | -105.7029
26.1097 | -105.7046
26.1134 | -105.709
26.1171 | -105.7126
26.1208 | -105.7174
26.1245 | -105.7241
26.1282 | -105.7294
26.1319 | -105.735
26.1355 | -105.7404
26.1392 | -105.7441
26.1429 | -105.7482
26.1466 | -105.7523
26.1503 | -105.7546
26.154 | -105.7559
26.1577 | -105.7601
26.1614 | -105.7652
26.1651 | -105.7677
26.1687 | -105.7706
26.1724 | -105.773
26.1761 | -105.7764
26.1798 | -105.7805
26.1835 | -105.7824
26.1872 | -105.7837
26.1909 | -105.7857
26.1946 | -105.7867
26.1982 | -105.7878
26.2019 | -105.7883
26.2056 | -105.7887
26.2093 | -105.7878
26.213 | -105.7877
26.2167 | -105.7875
26.2204 | -105.7842
26.2241 | -105.7815
26.2278 | -105.7792
26.2314 | -105.7757
26.2351 | -105.7724
26.2388 | -105.7704
26.2425 | -105.766
26.2462 | -105.7628
26.2499 | -105.7587
26.2536 | -105.7553
26.2573 | -105.7521
26.261 | -105.748
26.2646 | -105.7442
26.2683 | -105.7402
26.272 | -105.7368
26.2757 | -105.7347
26.2794 | -105.7316
26.2831 | -105.7279
26.2868 | -105.7257
26.2905 | -105.7237
26.2941 | -105.7201
26.2978 | -105.7159
26.3015 | -105.713
26.3052 | -105.7098
26.3089 | -105.7076
26.3126 | -105.7063
26.3163 | -105.7063
26.32 | -105.7067
26.3237 | -105.7069
26.3273 | -105.7077
26.331 | -105.7086
26.3347 | -105.7104
26.3384 | -105.7128
26.3421 | -105.7146
26.3458 | -105.7177
26.3495 | -105.7207
26.3532 | -105.7242
26.3568 | -105.7291
26.3605 | -105.7332
26.3642 | -105.737
26.3679 | -105.7416
26.3716 | -105.7467
26.3753 | -105.7535
26.379 | -105.7606
26.3827 | -105.766
26.3864 | -105.771
26.39 | -105.778
26.3937 | -105.7862
26.3974 | -105.7935
26.4011 | -105.8012
26.4048 | -105.8083
26.4085 | -105.8146
26.4122 | -105.8207
26.4159 | -105.8293
26.4195 | -105.8375
26.4232 | -105.8442
26.4269 | -105.8504
26.4306 | -105.8553
26.4343 | -105.8603
26.438 | -105.8677
26.4417 | -105.8744
26.4454 | -105.8806
26.4491 | -105.8856
26.4527 | -105.8894
26.4564 | -105.8956
26.4601 | -105.9019
26.4638 | -105.9082
26.4675 | -105.9136
26.4712 | -105.9172
26.4749 | -105.9206
26.4786 | -105.9261
26.4823 | -105.9317
26.4859 | -105.9334
26.4896 | -105.9347
26.4933 | -105.9363
26.497 | -105.9361
26.5007 | -105.9378
26.5044 | -105.9375
26.5081 | -105.9364
26.5118 | -105.9362
26.5154 | -105.9361
26.5191 | -105.935
26.5228 | -105.9352
26.5265 | -105.937
26.5302 | -105.934
26.5339 | -105.9299
26.5376 | -105.9284
26.5413 | -105.9254
26.545 | -105.9212
26.5486 | -105.9158
26.5523 | -105.9114
26.556 | -105.9063
26.5597 | -105.9008
26.5634 | -105.8953
26.5671 | -105.8905
26.5708 | -105.8858
26.5745 | -105.882
26.5781 | -105.8774
26.5818 | -105.8739
26.5855 | -105.8701
26.5892 | -105.8653
26.5929 | -105.8614
26.5966 | -105.8565
26.6003 | -105.8514
26.604 | -105.846
26.6077 | -105.8415
26.6113 | -105.8343
26.615 | -105.827
26.6187 | -105.8224
26.6224 | -105.8168
26.6261 | -105.8117
26.6298 | -105.807
26.6335 | -105.8022
26.6372 | -105.798
26.6408 | -105.7932
26.6445 | -105.7895
26.6482 | -105.7865
26.6519 | -105.7824
26.6556 | -105.7786
26.6593 | -105.7765
26.663 | -105.7743
26.6667 | -105.7718
26.6704 | -105.77
26.674 | -105.7689
26.6777 | -105.7684
26.6814 | -105.7684
26.6851 | -105.7674
26.6888 | -105.7659
26.6925 | -105.7653
26.6962 | -105.7659
26.6999 | -105.7654
26.7035 | -105.7641
26.7072 | -105.7642
26.7109 | -105.7659
26.7146 | -105.7664
26.7183 | -105.7668
26.722 | -105.7672
26.7257 | -105.7688
26.7294 | -105.7697
26.7331 | -105.7689
26.7367 | -105.77
26.7404 | -105.7706
26.7441 | -105.7713
26.7478 | -105.7725
26.7515 | -105.7737
26.7552 | -105.7749
26.7589 | -105.7772
26.7626 | -105.7789
26.7663 | -105.78
26.7699 | -105.7803
26.7736 | -105.781
26.7773 | -105.7816
26.781 | -105.7803
26.7847 | -105.7794
26.7884 | -105.7795
26.7921 | -105.7786
26.7958 | -105.7778
26.7994 | -105.777
26.8031 | -105.7765
26.8068 | -105.7775
26.8105 | -105.7785
26.8142 | -105.7777
26.8179 | -105.7758
26.8216 | -105.7734
26.8253 | -105.77
26.829 | -105.7676
26.8326 | -105.7657
26.8363 | -105.7631
26.84 | -105.7608
26.8437 | -105.7587
26.8474 | -105.7561
26.8511 | -105.7544
26.8548 | -105.7517
26.8585 | -105.7495
26.8621 | -105.7492
26.8658 | -105.7487
26.8695 | -105.7483
26.8732 | -105.7482
26.8769 | -105.7464
26.8806 | -105.7456
26.8843 | -105.7461
26.888 | -105.746
26.8917 | -105.745
26.8953 | -105.7436
26.899 | -105.7439
26.9027 | -105.7438
26.9064 | -105.7438
26.9101 | -105.7453
26.9138 | -105.7465
26.9175 | -105.7467
26.9212 | -105.7464
26.9248 | -105.7472
26.9285 | -105.7481
26.9322 | -105.7508
26.9359 | -105.7535
26.9396 | -105.7562
26.9433 | -105.7578
26.947 | -105.759
26.9507 | -105.7603
26.9544 | -105.7624
26.958 | -105.764
26.9617 | -105.7653
26.9654 | -105.7694
26.9691 | -105.7723
26.9728 | -105.7753
26.9765 | -105.7766
26.9802 | -105.7782
26.9839 | -105.7812
26.9876 | -105.7857
26.9912 | -105.791
26.9949 | -105.7944
26.9986 | -105.7946
27.0023 | -105.7952
27.006 | -105.7972
27.0097 | -105.7974
27.0134 | -105.7969
27.0171 | -105.7969
27.0207 | -105.7971
27.0244 | -105.7975
27.0281 | -105.7976
27.0318 | -105.7965
27.0355 | -105.7959
27.0392 | -105.7933
27.0429 | -105.792
27.0466 | -105.79
27.0503 | -105.7873
27.0539 | -105.7847
27.0576 | -105.7805
27.0613 | -105.7765
27.065 | -105.7738
27.0687 | -105.7714
27.0724 | -105.7669
27.0761 | -105.7621
27.0798 | -105.7581
27.0834 | -105.7545
27.0871 | -105.7492
27.0908 | -105.7416
27.0945 | -105.7363
27.0982 | -105.7324
27.1019 | -105.7266
27.1056 | -105.7185
27.1093 | -105.7116
27.113 | -105.7065
27.1166 | -105.6998
27.1203 | -105.6938
27.124 | -105.6877
27.1277 | -105.6805
27.1314 | -105.672
27.1351 | -105.6668
27.1388 | -105.66
27.1425 | -105.6534
27.1461 | -105.6466
27.1498 | -105.6389
27.1535 | -105.6316
27.1572 | -105.6258
27.1609 | -105.62
27.1646 | -105.6142
27.1683 | -105.6091
27.172 | -105.6047
27.1757 | -105.5991
27.1793 | -105.5947
27.183 | -105.5903
27.1867 | -105.5864
27.1904 | -105.5817
27.1941 | -105.5769
27.1978 | -105.5711
27.2015 | -105.5658
27.2052 | -105.5635
27.2089 | -105.5632
27.2125 | -105.5619
27.2162 | -105.5585
27.2199 | -105.5565
27.2236 | -105.5558
27.2273 | -105.5543
27.231 | -105.5526
27.2347 | -105.5508
27.2384 | -105.5506
27.242 | -105.5498
27.2457 | -105.5471
27.2494 | -105.546
27.2531 | -105.5466
27.2568 | -105.5475
27.2605 | -105.5487
27.2642 | -105.549
27.2679 | -105.5475
27.2716 | -105.547
27.2752 | -105.5482
27.2789 | -105.5503
27.2826 | -105.5505
27.2863 | -105.5513
27.29 | -105.5552
27.2937 | -105.5572
27.2974 | -105.5581
27.3011 | -105.5603
27.3047 | -105.5626
27.3084 | -105.5643
27.3121 | -105.5664
27.3158 | -105.5659
27.3195 | -105.567
27.3232 | -105.5681
27.3269 | -105.569
27.3306 | -105.5703
27.3343 | -105.5718
27.3379 | -105.5728
27.3416 | -105.5748
27.3453 | -105.5768
27.349 | -105.5772
27.3527 | -105.5799
27.3564 | -105.5821
27.3601 | -105.5856
27.3638 | -105.5897
27.3675 | -105.5908
27.3711 | -105.5901
27.3748 | -105.5915
27.3785 | -105.5924
27.3822 | -105.5946
27.3859 | -105.5975
27.3896 | -105.5995
27.3933 | -105.6013
27.397 | -105.6035
27.4006 | -105.6036
27.4043 | -105.6026
27.408 | -105.6017
27.4117 | -105.6003
27.4154 | -105.5996
27.4191 | -105.5994
27.4228 | -105.5984
27.4265 | -105.5957
27.4302 | -105.5937
27.4338 | -105.5931
27.4375 | -105.5915
27.4412 | -105.5895
27.4449 | -105.5868
27.4486 | -105.5843
27.4523 | -105.5803
27.456 | -105.5764
27.4597 | -105.5731
27.4633 | -105.5697
27.467 | -105.5666
27.4707 | -105.5643
27.4744 | -105.5629
27.4781 | -105.5607
27.4818 | -105.5572
27.4855 | -105.554
27.4892 | -105.5512
27.4929 | -105.5484
27.4965 | -105.544
27.5002 | -105.5398
27.5039 | -105.5351
27.5076 | -105.5296
27.5113 | -105.5244
27.515 | -105.5198
27.5187 | -105.5141
27.5224 | -105.509
27.526 | -105.5031
27.5297 | -105.4975
27.5334 | -105.4927
27.5371 | -105.4889
27.5408 | -105.4851
27.5445 | -105.4795
27.5482 | -105.473
27.5519 | -105.4688
27.5556 | -105.462
27.5592 | -105.4559
27.5629 | -105.4516
27.5666 | -105.4447
27.5703 | -105.4389
27.574 | -105.4342
27.5777 | -105.4295
27.5814 | -105.4231
27.5851 | -105.4168
27.5888 | -105.4132
27.5924 | -105.4071
27.5961 | -105.4005
27.5998 | -105.3956
27.6035 | -105.3921
27.6072 | -105.3875
27.6109 | -105.3831
27.6146 | -105.3773
27.6183 | -105.3731
27.6219 | -105.3704
27.6256 | -105.3676
27.6293 | -105.3656
27.633 | -105.3626
27.6367 | -105.3584
27.6404 | -105.3537
27.6441 | -105.3506
27.6478 | -105.3492
27.6515 | -105.3486
27.6551 | -105.349
27.6588 | -105.3481
27.6625 | -105.3468
27.6662 | -105.348
27.6699 | -105.3475
27.6736 | -105.3483
27.6773 | -105.348
27.681 | -105.3469
27.6846 | -105.3475
27.6883 | -105.3481
27.692 | -105.3478
27.6957 | -105.3481
27.6994 | -105.3482
27.7031 | -105.3488
27.7068 | -105.351
27.7105 | -105.3521
27.7142 | -105.3533
27.7178 | -105.3561
27.7215 | -105.3592
27.7252 | -105.3611
27.7289 | -105.3636
27.7326 | -105.3656
27.7363 | -105.3679
27.74 | -105.3713
27.7437 | -105.3731
27.7473 | -105.3755
27.751 | -105.378
27.7547 | -105.38
27.7584 | -105.3829
27.7621 | -105.3871
27.7658 | -105.3899
27.7695 | -105.3955
27.7732 | -105.4004
27.7769 | -105.4028
27.7805 | -105.4046
27.7842 | -105.4068
27.7879 | -105.4096
27.7916 | -105.4128
27.7953 | -105.4162
27.799 | -105.4189
27.8027 | -105.4223
27.8064 | -105.4252
27.8101 | -105.4268
27.8137 | -105.4303
27.8174 | -105.4346
27.8211 | -105.4376
27.8248 | -105.4391
27.8285 | -105.4399
27.8322 | -105.4412
27.8359 | -105.4436
27.8396 | -105.4452
27.8432 | -105.4465
27.8469 | -105.4491
27.8506 | -105.4506
27.8543 | -105.4506
27.858 | -105.4506
27.8617 | -105.449
27.8654 | -105.4474
27.8691 | -105.4459
27.8728 | -105.4453
27.8764 | -105.4448
27.8801 | -105.4446
27.8838 | -105.444
27.8875 | -105.4444
27.8912 | -105.4439
27.8949 | -105.4422
27.8986 | -105.441
27.9023 | -105.4403
27.9059 | -105.4401
27.9096 | -105.4399
27.9133 | -105.4398
27.917 | -105.4408
27.9207 | -105.4403
27.9244 | -105.4374
27.9281 | -105.4307
27.9318 | -105.4248
27.9355 | -105.422
27.9391 | -105.4215
27.9428 | -105.4197
27.9465 | -105.4201
27.9502 | -105.4186
27.9539 | -105.4145
27.9576 | -105.4128
27.9613 | -105.4113
27.965 | -105.4098
27.9686 | -105.408
27.9723 | -105.4073
27.976 | -105.408
27.9797 | -105.4059
27.9834 | -105.4042
27.9871 | -105.4024
27.9908 | -105.3992
27.9945 | -105.3953
27.9982 | -105.3923
28.0018 | -105.3902