## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 216 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 214
Hood | 02-92B/P | 153
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 992.4 | hPa |  | PRC PRO 01
Humidity | 43.5 | %/Hr
Internal temperature of calibration bench | 22.81 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 997 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.1 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 1.99 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.8 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.03 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3134 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 12/04/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 16/04/2018
Engineer | Anthony HUE | Approval |  | 23/04/2018
CEO | Robert GLEMOT | Approval |  | 23/04/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5046 | 57.4143
0.012209 | -36.095 | 38.2683
0.018313 | -35.0343 | 28.1528
0.024417 | -34.5732 | 21.8582
0.030521 | -34.4618 | 18.0753
0.036626 | -34.3175 | 15.0185
0.04273 | -34.2241 | 12.6707
0.048834 | -34.1822 | 11.143
0.054938 | -34.1201 | 9.9794
0.061043 | -34.0593 | 8.8892
0.067147 | -34.0497 | 8.0634
0.073251 | -34.0519 | 7.3012
0.079355 | -34.0715 | 6.6397
0.08546 | -34.0867 | 6.0466
0.091564 | -34.0837 | 5.6079
0.097668 | -34.0675 | 5.3637
0.10377 | -34.0462 | 5.1165
0.10988 | -34.0297 | 4.8365
0.11598 | -34.0213 | 4.538
0.12209 | -34.024 | 4.248
0.12819 | -34.0351 | 3.9604
0.13429 | -34.0473 | 3.6706
0.1404 | -34.0508 | 3.3923
0.1465 | -34.0531 | 3.1339
0.15261 | -34.0506 | 2.9129
0.15871 | -34.0445 | 2.7493
0.16482 | -34.0369 | 2.6452
0.17092 | -34.0351 | 2.6045
0.17702 | -34.0334 | 2.5312
0.18313 | -34.0328 | 2.4528
0.18923 | -34.0327 | 2.3446
0.19534 | -34.0317 | 2.2191
0.20144 | -34.0306 | 2.0782
0.20754 | -34.028 | 1.9576
0.21365 | -34.0273 | 1.843
0.21975 | -34.0271 | 1.7444
0.22586 | -34.0285 | 1.656
0.23196 | -34.0312 | 1.5819
0.23807 | -34.0356 | 1.5169
0.24417 | -34.0399 | 1.4554
0.25027 | -34.0441 | 1.3904
0.25638 | -34.0468 | 1.3262
0.26248 | -34.0489 | 1.2575
0.26859 | -34.0503 | 1.1909
0.27469 | -34.0512 | 1.1193
0.2808 | -34.0527 | 1.0467
0.2869 | -34.0539 | 0.97638
0.293 | -34.0552 | 0.90998
0.29911 | -34.0564 | 0.85037
0.30521 | -34.0576 | 0.7915
0.31132 | -34.0586 | 0.75225
0.31742 | -34.0601 | 0.70925
0.32353 | -34.0607 | 0.67381
0.32963 | -34.0618 | 0.64407
0.33573 | -34.0631 | 0.62293
0.34184 | -34.0648 | 0.60078
0.34794 | -34.0673 | 0.57608
0.35405 | -34.0698 | 0.55126
0.36015 | -34.0727 | 0.525
0.36626 | -34.0751 | 0.49669
0.37236 | -34.0775 | 0.45921
0.37846 | -34.0801 | 0.41434
0.38457 | -34.0813 | 0.35664
0.39067 | -34.0827 | 0.30943
0.39678 | -34.0837 | 0.257
0.40288 | -34.0845 | 0.20239
0.40899 | -34.0848 | 0.14981
0.41509 | -34.0844 | 0.098414
0.42119 | -34.0845 | 0.054357
0.4273 | -34.0847 | 0.0072758
0.4334 | -34.0849 | -0.039595
0.43951 | -34.0854 | -0.083645
0.44561 | -34.0865 | -0.12418
0.45172 | -34.0882 | -0.15658
0.45782 | -34.0892 | -0.19435
0.46392 | -34.0907 | -0.23129
0.47003 | -34.092 | -0.26439
0.47613 | -34.0931 | -0.29903
0.48224 | -34.0939 | -0.33423
0.48834 | -34.0933 | -0.36928
0.49445 | -34.0934 | -0.40484
0.50055 | -34.0931 | -0.43735
0.50665 | -34.0927 | -0.47176
0.51276 | -34.092 | -0.50612
0.51886 | -34.0912 | -0.54037
0.52497 | -34.0903 | -0.56481
0.53107 | -34.0897 | -0.59453
0.53717 | -34.0893 | -0.623
0.54328 | -34.0887 | -0.64617
0.54938 | -34.0887 | -0.66903
0.55549 | -34.0889 | -0.69027
0.56159 | -34.0902 | -0.70243
0.5677 | -34.0907 | -0.71755
0.5738 | -34.0919 | -0.73451
0.5799 | -34.0929 | -0.75
0.58601 | -34.0941 | -0.76072
0.59211 | -34.0959 | -0.76784
0.59822 | -34.0975 | -0.77805
0.60432 | -34.0992 | -0.79079
0.61043 | -34.1006 | -0.79951
0.61653 | -34.1021 | -0.81052
0.62263 | -34.1038 | -0.82069
0.62874 | -34.1047 | -0.83227
0.63484 | -34.1058 | -0.84236
0.64095 | -34.1065 | -0.8534
0.64705 | -34.1071 | -0.86409
0.65316 | -34.1075 | -0.87738
0.65926 | -34.1078 | -0.89394
0.66536 | -34.1069 | -0.91221
0.67147 | -34.1067 | -0.92627
0.67757 | -34.1066 | -0.94142
0.68368 | -34.106 | -0.96111
0.68978 | -34.1054 | -0.9792
0.69589 | -34.1047 | -0.99897
0.70199 | -34.1039 | -1.0223
0.70809 | -34.1032 | -1.0421
0.7142 | -34.1027 | -1.0632
0.7203 | -34.1017 | -1.0831
0.72641 | -34.1012 | -1.1061
0.73251 | -34.1007 | -1.1268
0.73862 | -34.1011 | -1.149
0.74472 | -34.1009 | -1.1696
0.75082 | -34.1008 | -1.1921
0.75693 | -34.1012 | -1.2145
0.76303 | -34.1019 | -1.237
0.76914 | -34.1038 | -1.258
0.77524 | -34.1049 | -1.2783
0.78135 | -34.1062 | -1.299
0.78745 | -34.1077 | -1.3196
0.79355 | -34.1091 | -1.3401
0.79966 | -34.1109 | -1.3603
0.80576 | -34.1133 | -1.3809
0.81187 | -34.115 | -1.4033
0.81797 | -34.117 | -1.4253
0.82408 | -34.1188 | -1.4467
0.83018 | -34.1207 | -1.468
0.83628 | -34.1227 | -1.4875
0.84239 | -34.124 | -1.509
0.84849 | -34.1255 | -1.5316
0.8546 | -34.1271 | -1.5524
0.8607 | -34.1285 | -1.574
0.86681 | -34.1297 | -1.5939
0.87291 | -34.1308 | -1.6137
0.87901 | -34.1314 | -1.6325
0.88512 | -34.1323 | -1.6515
0.89122 | -34.1332 | -1.6706
0.89733 | -34.1337 | -1.6871
0.90343 | -34.1343 | -1.7054
0.90953 | -34.1346 | -1.7204
0.91564 | -34.1348 | -1.737
0.92174 | -34.1351 | -1.7527
0.92785 | -34.1354 | -1.768
0.93395 | -34.1356 | -1.7825
0.94006 | -34.136 | -1.796
0.94616 | -34.1358 | -1.8082
0.95226 | -34.1361 | -1.8215
0.95837 | -34.1364 | -1.8358
0.96447 | -34.1367 | -1.849
0.97058 | -34.1367 | -1.8628
0.97668 | -34.1367 | -1.8766
0.98279 | -34.1366 | -1.8895
0.98889 | -34.1365 | -1.9006
0.99499 | -34.1367 | -1.9142
1.0011 | -34.1367 | -1.9253
1.0072 | -34.1365 | -1.9372
1.0133 | -34.1363 | -1.9482
1.0194 | -34.1361 | -1.9614
1.0255 | -34.1359 | -1.9754
1.0316 | -34.1355 | -1.9876
1.0377 | -34.1353 | -2.0014
1.0438 | -34.1348 | -2.0144
1.0499 | -34.1342 | -2.0298
1.056 | -34.1337 | -2.0443
1.0621 | -34.1332 | -2.0589
1.0682 | -34.1329 | -2.0715
1.0743 | -34.1324 | -2.0871
1.0805 | -34.1321 | -2.1018
1.0866 | -34.1313 | -2.1172
1.0927 | -34.1306 | -2.1319
1.0988 | -34.1301 | -2.1471
1.1049 | -34.1294 | -2.1616
1.111 | -34.1288 | -2.1772
1.1171 | -34.1282 | -2.1921
1.1232 | -34.1279 | -2.2076
1.1293 | -34.1276 | -2.2219
1.1354 | -34.1274 | -2.2373
1.1415 | -34.1272 | -2.2531
1.1476 | -34.1267 | -2.269
1.1537 | -34.1263 | -2.284
1.1598 | -34.1265 | -2.2992
1.1659 | -34.1264 | -2.3129
1.172 | -34.1263 | -2.3272
1.1781 | -34.1261 | -2.3416
1.1842 | -34.1261 | -2.3562
1.1903 | -34.1266 | -2.3677
1.1964 | -34.1266 | -2.3815
1.2025 | -34.1268 | -2.3959
1.2086 | -34.1273 | -2.4087
1.2147 | -34.1274 | -2.4216
1.2209 | -34.1279 | -2.4333
1.227 | -34.1288 | -2.4445
1.2331 | -34.1293 | -2.4569
1.2392 | -34.1298 | -2.4692
1.2453 | -34.1303 | -2.4807
1.2514 | -34.1308 | -2.4911
1.2575 | -34.1316 | -2.5026
1.2636 | -34.1325 | -2.5121
1.2697 | -34.1332 | -2.5218
1.2758 | -34.134 | -2.5334
1.2819 | -34.1348 | -2.5439
1.288 | -34.1356 | -2.5551
1.2941 | -34.1366 | -2.5641
1.3002 | -34.1374 | -2.5728
1.3063 | -34.138 | -2.5823
1.3124 | -34.1388 | -2.5919
1.3185 | -34.1397 | -2.6015
1.3246 | -34.1405 | -2.6111
1.3307 | -34.1414 | -2.6202
1.3368 | -34.1423 | -2.6295
1.3429 | -34.1431 | -2.639
1.349 | -34.1437 | -2.6488
1.3551 | -34.1445 | -2.6587
1.3613 | -34.1452 | -2.6676
1.3674 | -34.1459 | -2.6781
1.3735 | -34.1464 | -2.6884
1.3796 | -34.147 | -2.698
1.3857 | -34.1474 | -2.7078
1.3918 | -34.1481 | -2.7184
1.3979 | -34.1488 | -2.728
1.404 | -34.1493 | -2.7398
1.4101 | -34.1498 | -2.7506
1.4162 | -34.1503 | -2.7617
1.4223 | -34.1509 | -2.7735
1.4284 | -34.1515 | -2.7839
1.4345 | -34.1519 | -2.7959
1.4406 | -34.1523 | -2.8097
1.4467 | -34.1528 | -2.8217
1.4528 | -34.1534 | -2.8335
1.4589 | -34.154 | -2.8458
1.465 | -34.1543 | -2.857
1.4711 | -34.1549 | -2.8687
1.4772 | -34.1555 | -2.8823
1.4833 | -34.156 | -2.8956
1.4894 | -34.1566 | -2.9082
1.4955 | -34.1573 | -2.9205
1.5016 | -34.1579 | -2.9334
1.5078 | -34.1584 | -2.9475
1.5139 | -34.1588 | -2.9609
1.52 | -34.1594 | -2.9738
1.5261 | -34.1598 | -2.9867
1.5322 | -34.1602 | -3.001
1.5383 | -34.1607 | -3.0139
1.5444 | -34.1613 | -3.0268
1.5505 | -34.1618 | -3.0404
1.5566 | -34.1624 | -3.0537
1.5627 | -34.163 | -3.0678
1.5688 | -34.1637 | -3.0816
1.5749 | -34.1641 | -3.0945
1.581 | -34.1645 | -3.108
1.5871 | -34.165 | -3.122
1.5932 | -34.1654 | -3.1352
1.5993 | -34.1659 | -3.1496
1.6054 | -34.1664 | -3.1631
1.6115 | -34.1669 | -3.1775
1.6176 | -34.1672 | -3.1903
1.6237 | -34.1675 | -3.2042
1.6298 | -34.1679 | -3.217
1.6359 | -34.1684 | -3.23
1.642 | -34.1689 | -3.2421
1.6482 | -34.1691 | -3.2554
1.6543 | -34.1696 | -3.2693
1.6604 | -34.17 | -3.2816
1.6665 | -34.1702 | -3.2943
1.6726 | -34.1705 | -3.3072
1.6787 | -34.1709 | -3.3199
1.6848 | -34.1709 | -3.3328
1.6909 | -34.1712 | -3.3458
1.697 | -34.1715 | -3.3583
1.7031 | -34.1717 | -3.3711
1.7092 | -34.1717 | -3.3835
1.7153 | -34.172 | -3.3958
1.7214 | -34.1719 | -3.4084
1.7275 | -34.1719 | -3.4213
1.7336 | -34.1719 | -3.4345
1.7397 | -34.1722 | -3.4472
1.7458 | -34.1722 | -3.4597
1.7519 | -34.1724 | -3.4722
1.758 | -34.1725 | -3.4845
1.7641 | -34.1727 | -3.4965
1.7702 | -34.1727 | -3.5097
1.7763 | -34.1728 | -3.5215
1.7824 | -34.1728 | -3.5343
1.7885 | -34.1729 | -3.5472
1.7947 | -34.1731 | -3.5595
1.8008 | -34.1732 | -3.5712
1.8069 | -34.1733 | -3.5831
1.813 | -34.1734 | -3.5957
1.8191 | -34.1735 | -3.6088
1.8252 | -34.1737 | -3.6219
1.8313 | -34.1739 | -3.6346
1.8374 | -34.1741 | -3.6468
1.8435 | -34.1742 | -3.6592
1.8496 | -34.1744 | -3.6726
1.8557 | -34.1746 | -3.6857
1.8618 | -34.175 | -3.6984
1.8679 | -34.1752 | -3.7109
1.874 | -34.1753 | -3.7233
1.8801 | -34.1756 | -3.7362
1.8862 | -34.1759 | -3.7486
1.8923 | -34.1761 | -3.7612
1.8984 | -34.1763 | -3.7737
1.9045 | -34.1765 | -3.7858
1.9106 | -34.1766 | -3.7984
1.9167 | -34.1767 | -3.8111
1.9228 | -34.1769 | -3.8234
1.9289 | -34.1771 | -3.8357
1.9351 | -34.1774 | -3.8471
1.9412 | -34.1777 | -3.8588
1.9473 | -34.178 | -3.8713
1.9534 | -34.1783 | -3.8833
1.9595 | -34.1786 | -3.8955
1.9656 | -34.1789 | -3.9083
1.9717 | -34.1792 | -3.9205
1.9778 | -34.1796 | -3.9324
1.9839 | -34.1799 | -3.9443
1.99 | -34.1802 | -3.9559
1.9961 | -34.1804 | -3.9685
2.0022 | -34.1807 | -3.9804
2.0083 | -34.1809 | -3.9921
2.0144 | -34.1812 | -4.004
2.0205 | -34.1815 | -4.0162
2.0266 | -34.182 | -4.0281
2.0327 | -34.1823 | -4.0401
2.0388 | -34.1828 | -4.0521
2.0449 | -34.183 | -4.064
2.051 | -34.1834 | -4.0758
2.0571 | -34.1838 | -4.0876
2.0632 | -34.1841 | -4.0995
2.0693 | -34.1843 | -4.1121
2.0754 | -34.1847 | -4.1238
2.0816 | -34.1851 | -4.1364
2.0877 | -34.1853 | -4.1484
2.0938 | -34.1858 | -4.1609
2.0999 | -34.186 | -4.173
2.106 | -34.1863 | -4.185
2.1121 | -34.1866 | -4.1974
2.1182 | -34.1869 | -4.2086
2.1243 | -34.1872 | -4.2198
2.1304 | -34.1875 | -4.2314
2.1365 | -34.1879 | -4.2435
2.1426 | -34.1882 | -4.2559
2.1487 | -34.1886 | -4.2675
2.1548 | -34.1891 | -4.2797
2.1609 | -34.1896 | -4.2916
2.167 | -34.19 | -4.3032
2.1731 | -34.1904 | -4.3156
2.1792 | -34.1907 | -4.327
2.1853 | -34.1912 | -4.3393
2.1914 | -34.1915 | -4.3511
2.1975 | -34.192 | -4.363
2.2036 | -34.1923 | -4.3755
2.2097 | -34.1926 | -4.3874
2.2158 | -34.193 | -4.3984
2.222 | -34.1934 | -4.4093
2.2281 | -34.1937 | -4.4211
2.2342 | -34.1941 | -4.4328
2.2403 | -34.1944 | -4.4438
2.2464 | -34.1947 | -4.4549
2.2525 | -34.195 | -4.4667
2.2586 | -34.1955 | -4.4777
2.2647 | -34.1958 | -4.4885
2.2708 | -34.1962 | -4.5001
2.2769 | -34.1966 | -4.5112
2.283 | -34.1969 | -4.5227
2.2891 | -34.1972 | -4.534
2.2952 | -34.1974 | -4.5449
2.3013 | -34.1977 | -4.5561
2.3074 | -34.1981 | -4.567
2.3135 | -34.1984 | -4.5787
2.3196 | -34.1987 | -4.5892
2.3257 | -34.1989 | -4.6003
2.3318 | -34.1992 | -4.6118
2.3379 | -34.1994 | -4.6229
2.344 | -34.1996 | -4.6346
2.3501 | -34.1999 | -4.6457
2.3562 | -34.2002 | -4.6564
2.3623 | -34.2004 | -4.6672
2.3685 | -34.2006 | -4.6781
2.3746 | -34.2008 | -4.6896
2.3807 | -34.2011 | -4.7005
2.3868 | -34.2013 | -4.7116
2.3929 | -34.2016 | -4.7226
2.399 | -34.2017 | -4.7335
2.4051 | -34.2021 | -4.7443
2.4112 | -34.2023 | -4.7553
2.4173 | -34.2026 | -4.7663
2.4234 | -34.2027 | -4.7775
2.4295 | -34.203 | -4.7886
2.4356 | -34.2032 | -4.7998
2.4417 | -34.2034 | -4.8113
2.4478 | -34.2036 | -4.8225
2.4539 | -34.2039 | -4.8331
2.46 | -34.204 | -4.8449
2.4661 | -34.2043 | -4.8563
2.4722 | -34.2045 | -4.8668
2.4783 | -34.2047 | -4.8776
2.4844 | -34.2048 | -4.889
2.4905 | -34.205 | -4.9003
2.4966 | -34.2052 | -4.9114
2.5027 | -34.2054 | -4.9222
2.5089 | -34.2056 | -4.9335
2.515 | -34.2057 | -4.9446
2.5211 | -34.2059 | -4.9557
2.5272 | -34.206 | -4.9669
2.5333 | -34.2062 | -4.9782
2.5394 | -34.2065 | -4.9893
2.5455 | -34.2066 | -5.001
2.5516 | -34.2069 | -5.0117
2.5577 | -34.2071 | -5.0235
2.5638 | -34.2073 | -5.0355
2.5699 | -34.2075 | -5.0472
2.576 | -34.2076 | -5.0591
2.5821 | -34.2078 | -5.0706
2.5882 | -34.208 | -5.0821
2.5943 | -34.2081 | -5.0933
2.6004 | -34.2084 | -5.1046
2.6065 | -34.2086 | -5.1153
2.6126 | -34.2088 | -5.1261
2.6187 | -34.209 | -5.1369
2.6248 | -34.2092 | -5.148
2.6309 | -34.2094 | -5.1597
2.637 | -34.2098 | -5.1717
2.6431 | -34.21 | -5.1825
2.6492 | -34.2101 | -5.1941
2.6554 | -34.2104 | -5.2053
2.6615 | -34.2106 | -5.2166
2.6676 | -34.2109 | -5.2284
2.6737 | -34.2112 | -5.2398
2.6798 | -34.2116 | -5.2512
2.6859 | -34.2118 | -5.2628
2.692 | -34.2121 | -5.274
2.6981 | -34.2123 | -5.2855
2.7042 | -34.2126 | -5.297
2.7103 | -34.2128 | -5.308
2.7164 | -34.2132 | -5.3193
2.7225 | -34.2135 | -5.3309
2.7286 | -34.2138 | -5.3422
2.7347 | -34.2141 | -5.3531
2.7408 | -34.2144 | -5.3636
2.7469 | -34.2147 | -5.3752
2.753 | -34.2151 | -5.3861
2.7591 | -34.2154 | -5.3977
2.7652 | -34.2157 | -5.4095
2.7713 | -34.2159 | -5.4206
2.7774 | -34.2163 | -5.4319
2.7835 | -34.2165 | -5.4432
2.7896 | -34.2168 | -5.454
2.7958 | -34.2171 | -5.4651
2.8019 | -34.2175 | -5.4759
2.808 | -34.2178 | -5.487
2.8141 | -34.2182 | -5.4985
2.8202 | -34.2186 | -5.5097
2.8263 | -34.2189 | -5.5209
2.8324 | -34.2193 | -5.5326
2.8385 | -34.2198 | -5.5438
2.8446 | -34.2202 | -5.5548
2.8507 | -34.2206 | -5.5661
2.8568 | -34.2211 | -5.5768
2.8629 | -34.2214 | -5.5878
2.869 | -34.2218 | -5.5993
2.8751 | -34.2221 | -5.6103
2.8812 | -34.2226 | -5.622
2.8873 | -34.223 | -5.6332
2.8934 | -34.2234 | -5.6446
2.8995 | -34.2238 | -5.656
2.9056 | -34.2241 | -5.6675
2.9117 | -34.2245 | -5.6789
2.9178 | -34.2249 | -5.6899
2.9239 | -34.2254 | -5.7018
2.93 | -34.2259 | -5.7136
2.9361 | -34.2262 | -5.7248
2.9423 | -34.2266 | -5.7358
2.9484 | -34.2271 | -5.7472
2.9545 | -34.2275 | -5.7585
2.9606 | -34.228 | -5.7703
2.9667 | -34.2284 | -5.7818
2.9728 | -34.2288 | -5.7934
2.9789 | -34.2293 | -5.8048
2.985 | -34.2298 | -5.8165
2.9911 | -34.2302 | -5.8281
2.9972 | -34.2308 | -5.8396
3.0033 | -34.2312 | -5.8514
3.0094 | -34.2317 | -5.8628
3.0155 | -34.2321 | -5.8741
3.0216 | -34.2326 | -5.8857
3.0277 | -34.233 | -5.8977
3.0338 | -34.2334 | -5.9094
3.0399 | -34.2339 | -5.9211
3.046 | -34.2343 | -5.933
3.0521 | -34.2347 | -5.9444
3.0582 | -34.235 | -5.9568
3.0643 | -34.2354 | -5.9686
3.0704 | -34.2358 | -5.9808
3.0765 | -34.2362 | -5.9922
3.0827 | -34.2365 | -6.0038
3.0888 | -34.2369 | -6.0154
3.0949 | -34.2374 | -6.0271
3.101 | -34.2378 | -6.0389
3.1071 | -34.2382 | -6.051
3.1132 | -34.2386 | -6.0631
3.1193 | -34.239 | -6.0749
3.1254 | -34.2395 | -6.0865
3.1315 | -34.2399 | -6.0988
3.1376 | -34.2402 | -6.1105
3.1437 | -34.2407 | -6.1224
3.1498 | -34.2411 | -6.1348
3.1559 | -34.2415 | -6.1465
3.162 | -34.2418 | -6.159
3.1681 | -34.2421 | -6.1712
3.1742 | -34.2425 | -6.1832
3.1803 | -34.2429 | -6.1951
3.1864 | -34.2433 | -6.2068
3.1925 | -34.2437 | -6.2187
3.1986 | -34.244 | -6.2309
3.2047 | -34.2443 | -6.2433
3.2108 | -34.2447 | -6.2548
3.2169 | -34.245 | -6.2668
3.223 | -34.2453 | -6.2789
3.2292 | -34.2457 | -6.291
3.2353 | -34.2459 | -6.3022
3.2414 | -34.2462 | -6.314
3.2475 | -34.2465 | -6.3257
3.2536 | -34.2469 | -6.3375
3.2597 | -34.2472 | -6.3488
3.2658 | -34.2475 | -6.3601
3.2719 | -34.2477 | -6.3715
3.278 | -34.2481 | -6.3834
3.2841 | -34.2484 | -6.3948
3.2902 | -34.2487 | -6.4066
3.2963 | -34.249 | -6.4184
3.3024 | -34.2492 | -6.4299
3.3085 | -34.2495 | -6.4417
3.3146 | -34.2497 | -6.4532
3.3207 | -34.2501 | -6.465
3.3268 | -34.2503 | -6.4768
3.3329 | -34.2506 | -6.4879
3.339 | -34.2508 | -6.4991
3.3451 | -34.2511 | -6.5106
3.3512 | -34.2513 | -6.5217
3.3573 | -34.2516 | -6.5333
3.3634 | -34.2518 | -6.5447
3.3696 | -34.2521 | -6.5564
3.3757 | -34.2523 | -6.5675
3.3818 | -34.2526 | -6.5787
3.3879 | -34.2528 | -6.5902
3.394 | -34.253 | -6.6018
3.4001 | -34.2533 | -6.613
3.4062 | -34.2534 | -6.6243
3.4123 | -34.2537 | -6.6352
3.4184 | -34.2539 | -6.6468
3.4245 | -34.2542 | -6.6579
3.4306 | -34.2543 | -6.669
3.4367 | -34.2546 | -6.6799
3.4428 | -34.2548 | -6.6906
3.4489 | -34.255 | -6.7019
3.455 | -34.2552 | -6.713
3.4611 | -34.2555 | -6.7246
3.4672 | -34.2557 | -6.736
3.4733 | -34.2559 | -6.7474
3.4794 | -34.2561 | -6.7585
3.4855 | -34.2563 | -6.7692
3.4916 | -34.2566 | -6.78
3.4977 | -34.2568 | -6.7916
3.5038 | -34.257 | -6.8028
3.5099 | -34.2571 | -6.8134
3.5161 | -34.2574 | -6.8241
3.5222 | -34.2575 | -6.8348
3.5283 | -34.2577 | -6.8459
3.5344 | -34.2579 | -6.857
3.5405 | -34.2581 | -6.8679
3.5466 | -34.2583 | -6.8792
3.5527 | -34.2585 | -6.8902
3.5588 | -34.2587 | -6.9012
3.5649 | -34.2588 | -6.9121
3.571 | -34.259 | -6.9231
3.5771 | -34.2592 | -6.9345
3.5832 | -34.2594 | -6.9447
3.5893 | -34.2596 | -6.9557
3.5954 | -34.2599 | -6.9668
3.6015 | -34.2601 | -6.978
3.6076 | -34.2603 | -6.9889
3.6137 | -34.2605 | -6.9993
3.6198 | -34.2609 | -7.0099
3.6259 | -34.261 | -7.0211
3.632 | -34.2613 | -7.032
3.6381 | -34.2615 | -7.0431
3.6442 | -34.2617 | -7.054
3.6503 | -34.262 | -7.0645
3.6565 | -34.2622 | -7.0756
3.6626 | -34.2624 | -7.0865
3.6687 | -34.2627 | -7.0975
3.6748 | -34.263 | -7.1083
3.6809 | -34.2633 | -7.1192
3.687 | -34.2635 | -7.1305
3.6931 | -34.2638 | -7.141
3.6992 | -34.2641 | -7.152
3.7053 | -34.2643 | -7.1628
3.7114 | -34.2645 | -7.1738
3.7175 | -34.2648 | -7.1848
3.7236 | -34.2651 | -7.1952
3.7297 | -34.2654 | -7.2063
3.7358 | -34.2657 | -7.2175
3.7419 | -34.2659 | -7.2286
3.748 | -34.2662 | -7.2395
3.7541 | -34.2664 | -7.2504
3.7602 | -34.2668 | -7.2612
3.7663 | -34.267 | -7.272
3.7724 | -34.2673 | -7.2829
3.7785 | -34.2677 | -7.2936
3.7846 | -34.2679 | -7.3044
3.7907 | -34.2683 | -7.3156
3.7969 | -34.2686 | -7.3267
3.803 | -34.2689 | -7.3372
3.8091 | -34.2692 | -7.348
3.8152 | -34.2695 | -7.3589
3.8213 | -34.2698 | -7.37
3.8274 | -34.2701 | -7.3806
3.8335 | -34.2705 | -7.3917
3.8396 | -34.2709 | -7.4026
3.8457 | -34.2712 | -7.4136
3.8518 | -34.2716 | -7.4245
3.8579 | -34.272 | -7.4353
3.864 | -34.2723 | -7.4465
3.8701 | -34.2727 | -7.4578
3.8762 | -34.273 | -7.4685
3.8823 | -34.2733 | -7.4796
3.8884 | -34.2737 | -7.4906
3.8945 | -34.2741 | -7.5018
3.9006 | -34.2745 | -7.5131
3.9067 | -34.2748 | -7.5245
3.9128 | -34.2752 | -7.5356
3.9189 | -34.2755 | -7.5463
3.925 | -34.2758 | -7.5574
3.9311 | -34.2761 | -7.568
3.9372 | -34.2765 | -7.5791
3.9434 | -34.2769 | -7.5902
3.9495 | -34.2773 | -7.6013
3.9556 | -34.2777 | -7.6121
3.9617 | -34.278 | -7.6227
3.9678 | -34.2784 | -7.6334
3.9739 | -34.2788 | -7.6444
3.98 | -34.2791 | -7.6555
3.9861 | -34.2794 | -7.6664
3.9922 | -34.2797 | -7.6774
3.9983 | -34.2801 | -7.688
4.0044 | -34.2805 | -7.6986
4.0105 | -34.2808 | -7.7096
4.0166 | -34.2812 | -7.7206
4.0227 | -34.2816 | -7.7315
4.0288 | -34.2819 | -7.743
4.0349 | -34.2822 | -7.7537
4.041 | -34.2825 | -7.7646
4.0471 | -34.2828 | -7.7756
4.0532 | -34.2832 | -7.7863
4.0593 | -34.2836 | -7.7974
4.0654 | -34.2839 | -7.8081
4.0715 | -34.2843 | -7.8195
4.0776 | -34.2846 | -7.8304
4.0838 | -34.285 | -7.8418
4.0899 | -34.2854 | -7.8525
4.096 | -34.2857 | -7.8633
4.1021 | -34.286 | -7.8741
4.1082 | -34.2863 | -7.8848
4.1143 | -34.2866 | -7.8954
4.1204 | -34.287 | -7.9061
4.1265 | -34.2874 | -7.9167
4.1326 | -34.2877 | -7.9273
4.1387 | -34.2881 | -7.938
4.1448 | -34.2884 | -7.949
4.1509 | -34.2887 | -7.9599
4.157 | -34.289 | -7.9711
4.1631 | -34.2894 | -7.9822
4.1692 | -34.2897 | -7.9929
4.1753 | -34.2901 | -8.0038
4.1814 | -34.2903 | -8.0143
4.1875 | -34.2907 | -8.0254
4.1936 | -34.291 | -8.0364
4.1997 | -34.2913 | -8.0469
4.2058 | -34.2916 | -8.058
4.2119 | -34.2919 | -8.0687
4.218 | -34.2922 | -8.0793
4.2241 | -34.2925 | -8.0899
4.2303 | -34.2929 | -8.1006
4.2364 | -34.2932 | -8.1113
4.2425 | -34.2934 | -8.1221
4.2486 | -34.2937 | -8.1331
4.2547 | -34.2939 | -8.1437
4.2608 | -34.2942 | -8.1546
4.2669 | -34.2945 | -8.165
4.273 | -34.2948 | -8.176
4.2791 | -34.2951 | -8.1868
4.2852 | -34.2954 | -8.1977
4.2913 | -34.2956 | -8.2082
4.2974 | -34.2959 | -8.2192
4.3035 | -34.2962 | -8.2302
4.3096 | -34.2965 | -8.2412
4.3157 | -34.2968 | -8.2518
4.3218 | -34.2971 | -8.2624
4.3279 | -34.2974 | -8.2732
4.334 | -34.2976 | -8.2839
4.3401 | -34.2979 | -8.2942
4.3462 | -34.2982 | -8.305
4.3523 | -34.2986 | -8.3154
4.3584 | -34.2988 | -8.3257
4.3645 | -34.299 | -8.3363
4.3707 | -34.2993 | -8.3467
4.3768 | -34.2996 | -8.3573
4.3829 | -34.2999 | -8.3684
4.389 | -34.3002 | -8.379
4.3951 | -34.3004 | -8.3895
4.4012 | -34.3007 | -8.4005
4.4073 | -34.3009 | -8.411
4.4134 | -34.3012 | -8.4215
4.4195 | -34.3015 | -8.4322
4.4256 | -34.3019 | -8.4426
4.4317 | -34.3021 | -8.4532
4.4378 | -34.3024 | -8.4635
4.4439 | -34.3026 | -8.4738
4.45 | -34.3028 | -8.4844
4.4561 | -34.3031 | -8.495
4.4622 | -34.3033 | -8.5057
4.4683 | -34.3036 | -8.5162
4.4744 | -34.3039 | -8.5271
4.4805 | -34.3041 | -8.5373
4.4866 | -34.3044 | -8.5479
4.4927 | -34.3047 | -8.5587
4.4988 | -34.3049 | -8.5692
4.5049 | -34.3052 | -8.5794
4.511 | -34.3055 | -8.5897
4.5172 | -34.3058 | -8.6002
4.5233 | -34.306 | -8.6112
4.5294 | -34.3063 | -8.6216
4.5355 | -34.3066 | -8.6322
4.5416 | -34.3068 | -8.6427
4.5477 | -34.3071 | -8.6532
4.5538 | -34.3074 | -8.6633
4.5599 | -34.3077 | -8.6738
4.566 | -34.3079 | -8.6844
4.5721 | -34.3082 | -8.6949
4.5782 | -34.3085 | -8.7057
4.5843 | -34.3088 | -8.7163
4.5904 | -34.3091 | -8.7269
4.5965 | -34.3094 | -8.7373
4.6026 | -34.3096 | -8.7483
4.6087 | -34.3099 | -8.7587
4.6148 | -34.3101 | -8.7695
4.6209 | -34.3104 | -8.78
4.627 | -34.3108 | -8.7908
4.6331 | -34.311 | -8.8012
4.6392 | -34.3112 | -8.8118
4.6453 | -34.3115 | -8.8223
4.6514 | -34.3118 | -8.8329
4.6576 | -34.3121 | -8.8439
4.6637 | -34.3123 | -8.854
4.6698 | -34.3125 | -8.8644
4.6759 | -34.3128 | -8.8754
4.682 | -34.3131 | -8.8858
4.6881 | -34.3133 | -8.8963
4.6942 | -34.3136 | -8.9068
4.7003 | -34.3138 | -8.9174
4.7064 | -34.3141 | -8.9277
4.7125 | -34.3144 | -8.9386
4.7186 | -34.3147 | -8.9489
4.7247 | -34.315 | -8.9595
4.7308 | -34.3152 | -8.97
4.7369 | -34.3155 | -8.9807
4.743 | -34.3158 | -8.9916
4.7491 | -34.3161 | -9.0024
4.7552 | -34.3163 | -9.0131
4.7613 | -34.3166 | -9.0238
4.7674 | -34.3168 | -9.0342
4.7735 | -34.3171 | -9.0452
4.7796 | -34.3174 | -9.056
4.7857 | -34.3177 | -9.0668
4.7918 | -34.318 | -9.0772
4.7979 | -34.3183 | -9.0877
4.8041 | -34.3185 | -9.0983
4.8102 | -34.3188 | -9.1086
4.8163 | -34.3191 | -9.1191
4.8224 | -34.3194 | -9.1297
4.8285 | -34.3197 | -9.14
4.8346 | -34.32 | -9.1504
4.8407 | -34.3202 | -9.1611
4.8468 | -34.3206 | -9.172
4.8529 | -34.3208 | -9.1828
4.859 | -34.3211 | -9.1937
4.8651 | -34.3214 | -9.2041
4.8712 | -34.3217 | -9.2145
4.8773 | -34.3219 | -9.2251
4.8834 | -34.3223 | -9.2358
4.8895 | -34.3225 | -9.2462
4.8956 | -34.3228 | -9.2568
4.9017 | -34.323 | -9.2672
4.9078 | -34.3233 | -9.2781
4.9139 | -34.3236 | -9.2883
4.92 | -34.324 | -9.2986
4.9261 | -34.3243 | -9.3088
4.9322 | -34.3246 | -9.3195
4.9383 | -34.3249 | -9.33
4.9445 | -34.3252 | -9.3407
4.9506 | -34.3254 | -9.3511
4.9567 | -34.3257 | -9.3614
4.9628 | -34.326 | -9.3719
4.9689 | -34.3263 | -9.3829
4.975 | -34.3266 | -9.3938
4.9811 | -34.3269 | -9.4041
4.9872 | -34.3272 | -9.4145
4.9933 | -34.3276 | -9.4254
4.9994 | -34.3279 | -9.4364
5.0055 | -34.3282 | -9.447
5.0116 | -34.3285 | -9.458
5.0177 | -34.3289 | -9.4689
5.0238 | -34.3292 | -9.4798
5.0299 | -34.3296 | -9.4907
5.036 | -34.3299 | -9.5014
5.0421 | -34.3302 | -9.5119
5.0482 | -34.3305 | -9.5226
5.0543 | -34.3308 | -9.5329
5.0604 | -34.3311 | -9.5435
5.0665 | -34.3315 | -9.5539
5.0726 | -34.3318 | -9.5648
5.0787 | -34.3321 | -9.5757
5.0848 | -34.3325 | -9.5864
5.091 | -34.3328 | -9.5969
5.0971 | -34.3331 | -9.6072
5.1032 | -34.3334 | -9.6181
5.1093 | -34.3338 | -9.6288
5.1154 | -34.334 | -9.6392
5.1215 | -34.3343 | -9.6496
5.1276 | -34.3347 | -9.66
5.1337 | -34.335 | -9.6704
5.1398 | -34.3353 | -9.6808
5.1459 | -34.3357 | -9.6916
5.152 | -34.336 | -9.7024
5.1581 | -34.3363 | -9.7133
5.1642 | -34.3367 | -9.724
5.1703 | -34.337 | -9.7346
5.1764 | -34.3374 | -9.7451
5.1825 | -34.3377 | -9.7555
5.1886 | -34.338 | -9.7658
5.1947 | -34.3383 | -9.7765
5.2008 | -34.3387 | -9.7873
5.2069 | -34.339 | -9.7981
5.213 | -34.3393 | -9.8086
5.2191 | -34.3397 | -9.8193
5.2252 | -34.34 | -9.8299
5.2314 | -34.3403 | -9.8408
5.2375 | -34.3406 | -9.8512
5.2436 | -34.3409 | -9.8619
5.2497 | -34.3413 | -9.8723
5.2558 | -34.3417 | -9.8833
5.2619 | -34.342 | -9.8942
5.268 | -34.3424 | -9.905
5.2741 | -34.3427 | -9.9158
5.2802 | -34.3431 | -9.9268
5.2863 | -34.3434 | -9.9374
5.2924 | -34.3438 | -9.9482
5.2985 | -34.3441 | -9.959
5.3046 | -34.3445 | -9.9697
5.3107 | -34.3448 | -9.9806
5.3168 | -34.3452 | -9.9914
5.3229 | -34.3455 | -10.0021
5.329 | -34.3458 | -10.0128
5.3351 | -34.3462 | -10.0232
5.3412 | -34.3465 | -10.0339
5.3473 | -34.3469 | -10.0447
5.3534 | -34.3473 | -10.0555
5.3595 | -34.3476 | -10.0663
5.3656 | -34.348 | -10.077
5.3717 | -34.3483 | -10.0881
5.3779 | -34.3486 | -10.0991
5.384 | -34.349 | -10.1098
5.3901 | -34.3494 | -10.1208
5.3962 | -34.3497 | -10.1313
5.4023 | -34.3501 | -10.1423
5.4084 | -34.3504 | -10.153
5.4145 | -34.3508 | -10.1635
5.4206 | -34.3511 | -10.1742
5.4267 | -34.3515 | -10.1853
5.4328 | -34.3518 | -10.1961
5.4389 | -34.3522 | -10.2066
5.445 | -34.3525 | -10.2177
5.4511 | -34.3529 | -10.2282
5.4572 | -34.3533 | -10.239
5.4633 | -34.3536 | -10.2498
5.4694 | -34.3539 | -10.2606
5.4755 | -34.3543 | -10.2711
5.4816 | -34.3546 | -10.2819
5.4877 | -34.355 | -10.2925
5.4938 | -34.3554 | -10.303
5.4999 | -34.3558 | -10.3136
5.506 | -34.3561 | -10.3243
5.5121 | -34.3565 | -10.335
5.5183 | -34.3568 | -10.3456
5.5244 | -34.3572 | -10.3565
5.5305 | -34.3576 | -10.3672
5.5366 | -34.3579 | -10.3777
5.5427 | -34.3582 | -10.3886
5.5488 | -34.3586 | -10.3994
5.5549 | -34.3589 | -10.4102
5.561 | -34.3593 | -10.4212
5.5671 | -34.3596 | -10.4318
5.5732 | -34.36 | -10.4427
5.5793 | -34.3604 | -10.4539
5.5854 | -34.3608 | -10.4645
5.5915 | -34.3611 | -10.4756
5.5976 | -34.3614 | -10.4864
5.6037 | -34.3618 | -10.4973
5.6098 | -34.3621 | -10.508
5.6159 | -34.3625 | -10.5186
5.622 | -34.3628 | -10.5294
5.6281 | -34.3632 | -10.5403
5.6342 | -34.3636 | -10.5514
5.6403 | -34.3639 | -10.5622
5.6464 | -34.3643 | -10.5731
5.6525 | -34.3646 | -10.584
5.6586 | -34.365 | -10.5944
5.6648 | -34.3654 | -10.6052
5.6709 | -34.3658 | -10.6162
5.677 | -34.3662 | -10.627
5.6831 | -34.3665 | -10.6378
5.6892 | -34.3668 | -10.6488
5.6953 | -34.3672 | -10.6596
5.7014 | -34.3676 | -10.6705
5.7075 | -34.3679 | -10.6813
5.7136 | -34.3682 | -10.692
5.7197 | -34.3686 | -10.7028
5.7258 | -34.3689 | -10.7134
5.7319 | -34.3693 | -10.7242
5.738 | -34.3697 | -10.735
5.7441 | -34.3701 | -10.7457
5.7502 | -34.3705 | -10.7562
5.7563 | -34.3708 | -10.767
5.7624 | -34.3712 | -10.7779
5.7685 | -34.3716 | -10.7888
5.7746 | -34.3719 | -10.7998
5.7807 | -34.3722 | -10.8106
5.7868 | -34.3726 | -10.8216
5.7929 | -34.3729 | -10.8323
5.799 | -34.3733 | -10.843
5.8052 | -34.3736 | -10.8536
5.8113 | -34.374 | -10.8643
5.8174 | -34.3743 | -10.8749
5.8235 | -34.3747 | -10.8856
5.8296 | -34.375 | -10.8966
5.8357 | -34.3754 | -10.9073
5.8418 | -34.3757 | -10.9182
5.8479 | -34.376 | -10.9291
5.854 | -34.3764 | -10.94
5.8601 | -34.3768 | -10.9506
5.8662 | -34.3771 | -10.9615
5.8723 | -34.3775 | -10.9721
5.8784 | -34.3778 | -10.9828
5.8845 | -34.3782 | -10.9934
5.8906 | -34.3785 | -11.0043
5.8967 | -34.3788 | -11.0148
5.9028 | -34.3792 | -11.0254
5.9089 | -34.3797 | -11.036
5.915 | -34.38 | -11.0466
5.9211 | -34.3803 | -11.0573
5.9272 | -34.3807 | -11.0678
5.9333 | -34.381 | -11.0784
5.9394 | -34.3814 | -11.0891
5.9455 | -34.3818 | -11.0996
5.9517 | -34.3822 | -11.1102
5.9578 | -34.3826 | -11.121
5.9639 | -34.3829 | -11.1314
5.97 | -34.3833 | -11.1421
5.9761 | -34.3837 | -11.1528
5.9822 | -34.3841 | -11.1635
5.9883 | -34.3844 | -11.1742
5.9944 | -34.3848 | -11.185
6.0005 | -34.3852 | -11.1958
6.0066 | -34.3855 | -11.2063
6.0127 | -34.3859 | -11.2168
6.0188 | -34.3863 | -11.2277
6.0249 | -34.3866 | -11.2385
6.031 | -34.387 | -11.249
6.0371 | -34.3873 | -11.2597
6.0432 | -34.3877 | -11.2703
6.0493 | -34.3881 | -11.2807
6.0554 | -34.3885 | -11.2913
6.0615 | -34.3888 | -11.3019
6.0676 | -34.3892 | -11.3128
6.0737 | -34.3896 | -11.3236
6.0798 | -34.3899 | -11.3342
6.0859 | -34.3904 | -11.3448
6.0921 | -34.3907 | -11.3552
6.0982 | -34.3911 | -11.3658
6.1043 | -34.3915 | -11.3763
6.1104 | -34.3919 | -11.3872
6.1165 | -34.3922 | -11.3973
6.1226 | -34.3926 | -11.4079
6.1287 | -34.393 | -11.4185
6.1348 | -34.3934 | -11.4291
6.1409 | -34.3937 | -11.4396
6.147 | -34.3941 | -11.4502
6.1531 | -34.3944 | -11.4607
6.1592 | -34.3948 | -11.4712
6.1653 | -34.3951 | -11.4817
6.1714 | -34.3955 | -11.4923
6.1775 | -34.3959 | -11.503
6.1836 | -34.3962 | -11.5136
6.1897 | -34.3966 | -11.5242
6.1958 | -34.397 | -11.535
6.2019 | -34.3974 | -11.5457
6.208 | -34.3978 | -11.5563
6.2141 | -34.3981 | -11.5668
6.2202 | -34.3985 | -11.5775
6.2263 | -34.3989 | -11.588
6.2325 | -34.3993 | -11.5982
6.2386 | -34.3996 | -11.6085
6.2447 | -34.4 | -11.6193
6.2508 | -34.4004 | -11.6299
6.2569 | -34.4008 | -11.6403
6.263 | -34.4012 | -11.6509
6.2691 | -34.4015 | -11.6613
6.2752 | -34.4019 | -11.6719
6.2813 | -34.4022 | -11.6825
6.2874 | -34.4026 | -11.6929
6.2935 | -34.403 | -11.7033
6.2996 | -34.4033 | -11.7139
6.3057 | -34.4037 | -11.7244
6.3118 | -34.404 | -11.7351
6.3179 | -34.4044 | -11.7456
6.324 | -34.4048 | -11.756
6.3301 | -34.4051 | -11.7665
6.3362 | -34.4055 | -11.7769
6.3423 | -34.4059 | -11.7876
6.3484 | -34.4063 | -11.7982
6.3545 | -34.4066 | -11.8084
6.3606 | -34.407 | -11.8186
6.3667 | -34.4073 | -11.8289
6.3728 | -34.4077 | -11.8395
6.379 | -34.4081 | -11.85
6.3851 | -34.4085 | -11.8603
6.3912 | -34.4088 | -11.8707
6.3973 | -34.4091 | -11.8812
6.4034 | -34.4095 | -11.8914
6.4095 | -34.4099 | -11.9019
6.4156 | -34.4102 | -11.9124
6.4217 | -34.4106 | -11.9228
6.4278 | -34.4109 | -11.9334
6.4339 | -34.4113 | -11.9437
6.44 | -34.4117 | -11.9541
6.4461 | -34.4121 | -11.9648
6.4522 | -34.4124 | -11.975
6.4583 | -34.4128 | -11.9856
6.4644 | -34.4131 | -11.9961
6.4705 | -34.4135 | -12.0062
6.4766 | -34.4138 | -12.0167
6.4827 | -34.4142 | -12.0272
6.4888 | -34.4145 | -12.0374
6.4949 | -34.4149 | -12.0476
6.501 | -34.4153 | -12.0581
6.5071 | -34.4157 | -12.0683
6.5132 | -34.416 | -12.0787
6.5194 | -34.4164 | -12.0892
6.5255 | -34.4167 | -12.0994
6.5316 | -34.4171 | -12.1099
6.5377 | -34.4174 | -12.1203
6.5438 | -34.4178 | -12.1305
6.5499 | -34.4181 | -12.1408
6.556 | -34.4185 | -12.1511
6.5621 | -34.4188 | -12.1618
6.5682 | -34.4192 | -12.1722
6.5743 | -34.4196 | -12.1825
6.5804 | -34.42 | -12.193
6.5865 | -34.4203 | -12.2034
6.5926 | -34.4207 | -12.2137
6.5987 | -34.4211 | -12.2238
6.6048 | -34.4214 | -12.2342
6.6109 | -34.4218 | -12.2445
6.617 | -34.4221 | -12.255
6.6231 | -34.4225 | -12.2653
6.6292 | -34.4229 | -12.2757
6.6353 | -34.4232 | -12.2863
6.6414 | -34.4236 | -12.2966
6.6475 | -34.424 | -12.3069
6.6536 | -34.4243 | -12.3171
6.6597 | -34.4247 | -12.3272
6.6659 | -34.425 | -12.3376
6.672 | -34.4254 | -12.3478
6.6781 | -34.4258 | -12.3581
6.6842 | -34.4262 | -12.3683
6.6903 | -34.4266 | -12.3786
6.6964 | -34.4269 | -12.389
6.7025 | -34.4273 | -12.3992
6.7086 | -34.4276 | -12.4094
6.7147 | -34.428 | -12.4198
6.7208 | -34.4283 | -12.4302
6.7269 | -34.4287 | -12.4405
6.733 | -34.429 | -12.4506
6.7391 | -34.4294 | -12.4609
6.7452 | -34.4298 | -12.471
6.7513 | -34.4302 | -12.481
6.7574 | -34.4305 | -12.4914
6.7635 | -34.4309 | -12.5015
6.7696 | -34.4313 | -12.5116
6.7757 | -34.4317 | -12.5219
6.7818 | -34.432 | -12.5322
6.7879 | -34.4324 | -12.5423
6.794 | -34.4328 | -12.5524
6.8001 | -34.4331 | -12.5626
6.8063 | -34.4335 | -12.5729
6.8124 | -34.4338 | -12.5833
6.8185 | -34.4342 | -12.5937
6.8246 | -34.4345 | -12.6037
6.8307 | -34.4349 | -12.6139
6.8368 | -34.4353 | -12.6244
6.8429 | -34.4357 | -12.6346
6.849 | -34.436 | -12.6449
6.8551 | -34.4364 | -12.655
6.8612 | -34.4368 | -12.6652
6.8673 | -34.4371 | -12.6754
6.8734 | -34.4375 | -12.6858
6.8795 | -34.4379 | -12.6961
6.8856 | -34.4382 | -12.7063
6.8917 | -34.4386 | -12.7164
6.8978 | -34.4389 | -12.7264
6.9039 | -34.4393 | -12.7364
6.91 | -34.4397 | -12.7466
6.9161 | -34.4401 | -12.7566
6.9222 | -34.4404 | -12.7668
6.9283 | -34.4409 | -12.7767
6.9344 | -34.4412 | -12.787
6.9405 | -34.4416 | -12.7972
6.9466 | -34.442 | -12.8076
6.9528 | -34.4424 | -12.8179
6.9589 | -34.4428 | -12.8283
6.965 | -34.4432 | -12.8385
6.9711 | -34.4436 | -12.8488
6.9772 | -34.444 | -12.859
6.9833 | -34.4443 | -12.8692
6.9894 | -34.4447 | -12.8796
6.9955 | -34.4451 | -12.8896
7.0016 | -34.4455 | -12.8997
7.0077 | -34.4458 | -12.9099
7.0138 | -34.4462 | -12.92
7.0199 | -34.4466 | -12.9301
7.026 | -34.447 | -12.9402
7.0321 | -34.4474 | -12.9502
7.0382 | -34.4477 | -12.96
7.0443 | -34.4482 | -12.9704
7.0504 | -34.4485 | -12.9806
7.0565 | -34.4489 | -12.9908
7.0626 | -34.4493 | -13.0007
7.0687 | -34.4497 | -13.0105
7.0748 | -34.4501 | -13.0204
7.0809 | -34.4505 | -13.0305
7.087 | -34.4508 | -13.0407
7.0932 | -34.4512 | -13.0509
7.0993 | -34.4516 | -13.061
7.1054 | -34.452 | -13.0709
7.1115 | -34.4524 | -13.0809
7.1176 | -34.4527 | -13.0909
7.1237 | -34.4531 | -13.101
7.1298 | -34.4535 | -13.111
7.1359 | -34.4538 | -13.1211
7.142 | -34.4542 | -13.131
7.1481 | -34.4546 | -13.141
7.1542 | -34.455 | -13.1511
7.1603 | -34.4554 | -13.1612
7.1664 | -34.4558 | -13.1712
7.1725 | -34.4562 | -13.1814
7.1786 | -34.4566 | -13.1915
7.1847 | -34.457 | -13.2016
7.1908 | -34.4573 | -13.2118
7.1969 | -34.4577 | -13.2221
7.203 | -34.4581 | -13.232
7.2091 | -34.4585 | -13.242
7.2152 | -34.4589 | -13.2521
7.2213 | -34.4593 | -13.2621
7.2274 | -34.4597 | -13.2721
7.2335 | -34.46 | -13.2821
7.2397 | -34.4604 | -13.2922
7.2458 | -34.4609 | -13.302
7.2519 | -34.4612 | -13.3121
7.258 | -34.4616 | -13.3223
7.2641 | -34.462 | -13.3324
7.2702 | -34.4624 | -13.3426
7.2763 | -34.4628 | -13.3527
7.2824 | -34.4632 | -13.363
7.2885 | -34.4636 | -13.3731
7.2946 | -34.464 | -13.3831
7.3007 | -34.4643 | -13.3933
7.3068 | -34.4647 | -13.4036
7.3129 | -34.4651 | -13.4139
7.319 | -34.4655 | -13.424
7.3251 | -34.4659 | -13.4338
7.3312 | -34.4663 | -13.4439
7.3373 | -34.4667 | -13.4538
7.3434 | -34.4671 | -13.4637
7.3495 | -34.4675 | -13.4737
7.3556 | -34.4679 | -13.4838
7.3617 | -34.4683 | -13.4941
7.3678 | -34.4687 | -13.5041
7.3739 | -34.4691 | -13.5144
7.3801 | -34.4695 | -13.5247
7.3862 | -34.4699 | -13.5347
7.3923 | -34.4703 | -13.5449
7.3984 | -34.4707 | -13.5548
7.4045 | -34.4711 | -13.5648
7.4106 | -34.4715 | -13.5748
7.4167 | -34.4719 | -13.585
7.4228 | -34.4723 | -13.5951
7.4289 | -34.4727 | -13.6052
7.435 | -34.4731 | -13.615
7.4411 | -34.4735 | -13.6252
7.4472 | -34.4739 | -13.6353
7.4533 | -34.4743 | -13.6453
7.4594 | -34.4747 | -13.6556
7.4655 | -34.4751 | -13.6655
7.4716 | -34.4755 | -13.6757
7.4777 | -34.4759 | -13.6857
7.4838 | -34.4763 | -13.6957
7.4899 | -34.4767 | -13.7057
7.496 | -34.4771 | -13.7156
7.5021 | -34.4775 | -13.7257
7.5082 | -34.4779 | -13.7357
7.5143 | -34.4783 | -13.7458
7.5204 | -34.4787 | -13.756
7.5266 | -34.4791 | -13.766
7.5327 | -34.4795 | -13.7759
7.5388 | -34.4799 | -13.7863
7.5449 | -34.4803 | -13.7965
7.551 | -34.4807 | -13.8065
7.5571 | -34.4811 | -13.8167
7.5632 | -34.4816 | -13.8268
7.5693 | -34.482 | -13.8368
7.5754 | -34.4824 | -13.8467
7.5815 | -34.4828 | -13.8568
7.5876 | -34.4832 | -13.8671
7.5937 | -34.4836 | -13.8774
7.5998 | -34.484 | -13.8874
7.6059 | -34.4845 | -13.8977
7.612 | -34.4849 | -13.9077
7.6181 | -34.4853 | -13.9177
7.6242 | -34.4857 | -13.9281
7.6303 | -34.4861 | -13.9384
7.6364 | -34.4865 | -13.9485
7.6425 | -34.4869 | -13.9586
7.6486 | -34.4872 | -13.9688
7.6547 | -34.4877 | -13.9788
7.6608 | -34.4881 | -13.9889
7.667 | -34.4885 | -13.9991
7.6731 | -34.4889 | -14.0093
7.6792 | -34.4893 | -14.0196
7.6853 | -34.4897 | -14.0298
7.6914 | -34.4901 | -14.04
7.6975 | -34.4905 | -14.0502
7.7036 | -34.491 | -14.0604
7.7097 | -34.4914 | -14.0707
7.7158 | -34.4918 | -14.0809
7.7219 | -34.4922 | -14.0909
7.728 | -34.4926 | -14.1012
7.7341 | -34.493 | -14.1114
7.7402 | -34.4934 | -14.1216
7.7463 | -34.4939 | -14.1318
7.7524 | -34.4943 | -14.1422
7.7585 | -34.4947 | -14.1524
7.7646 | -34.4951 | -14.1625
7.7707 | -34.4955 | -14.1728
7.7768 | -34.4959 | -14.1831
7.7829 | -34.4963 | -14.1935
7.789 | -34.4967 | -14.204
7.7951 | -34.4971 | -14.2141
7.8012 | -34.4975 | -14.2243
7.8073 | -34.498 | -14.2347
7.8135 | -34.4984 | -14.2448
7.8196 | -34.4988 | -14.2553
7.8257 | -34.4992 | -14.2656
7.8318 | -34.4996 | -14.2759
7.8379 | -34.5 | -14.286
7.844 | -34.5004 | -14.2961
7.8501 | -34.5008 | -14.3065
7.8562 | -34.5013 | -14.3169
7.8623 | -34.5017 | -14.327
7.8684 | -34.5021 | -14.3373
7.8745 | -34.5025 | -14.3475
7.8806 | -34.5029 | -14.3579
7.8867 | -34.5033 | -14.3679
7.8928 | -34.5038 | -14.3779
7.8989 | -34.5042 | -14.3881
7.905 | -34.5046 | -14.3984
7.9111 | -34.505 | -14.4086
7.9172 | -34.5054 | -14.4188
7.9233 | -34.5058 | -14.4291
7.9294 | -34.5062 | -14.4394
7.9355 | -34.5066 | -14.4494
7.9416 | -34.507 | -14.4597
7.9477 | -34.5074 | -14.4698
7.9539 | -34.5078 | -14.4799
7.96 | -34.5083 | -14.4903
7.9661 | -34.5087 | -14.5006
7.9722 | -34.5091 | -14.5106
7.9783 | -34.5096 | -14.5211
7.9844 | -34.51 | -14.5313
7.9905 | -34.5104 | -14.5417
7.9966 | -34.5108 | -14.552
8.0027 | -34.5112 | -14.5623
8.0088 | -34.5116 | -14.5724
8.0149 | -34.512 | -14.5827
8.021 | -34.5124 | -14.593
8.0271 | -34.5129 | -14.6035
8.0332 | -34.5133 | -14.6137
8.0393 | -34.5137 | -14.6241
8.0454 | -34.5141 | -14.6345
8.0515 | -34.5146 | -14.6448
8.0576 | -34.515 | -14.6553
8.0637 | -34.5154 | -14.6655
8.0698 | -34.5158 | -14.6756
8.0759 | -34.5162 | -14.6857
8.082 | -34.5166 | -14.6958
8.0881 | -34.517 | -14.7062
8.0942 | -34.5175 | -14.7165
8.1004 | -34.5179 | -14.7269
8.1065 | -34.5183 | -14.737
8.1126 | -34.5187 | -14.7472
8.1187 | -34.5191 | -14.7576
8.1248 | -34.5195 | -14.7676
8.1309 | -34.52 | -14.7778
8.137 | -34.5204 | -14.7878
8.1431 | -34.5208 | -14.7983
8.1492 | -34.5212 | -14.8085
8.1553 | -34.5216 | -14.8188
8.1614 | -34.522 | -14.8292
8.1675 | -34.5224 | -14.8394
8.1736 | -34.5228 | -14.8499
8.1797 | -34.5232 | -14.8602
8.1858 | -34.5237 | -14.8705
8.1919 | -34.5241 | -14.8808
8.198 | -34.5245 | -14.8908
8.2041 | -34.5249 | -14.9011
8.2102 | -34.5254 | -14.9114
8.2163 | -34.5258 | -14.9214
8.2224 | -34.5262 | -14.9317
8.2285 | -34.5266 | -14.9417
8.2346 | -34.527 | -14.9518
8.2408 | -34.5274 | -14.9619
8.2469 | -34.5279 | -14.9722
8.253 | -34.5283 | -14.9823
8.2591 | -34.5287 | -14.9926
8.2652 | -34.5292 | -15.0029
8.2713 | -34.5296 | -15.0132
8.2774 | -34.53 | -15.0235
8.2835 | -34.5305 | -15.034
8.2896 | -34.5309 | -15.0441
8.2957 | -34.5313 | -15.0543
8.3018 | -34.5317 | -15.0646
8.3079 | -34.5322 | -15.0746
8.314 | -34.5326 | -15.0849
8.3201 | -34.533 | -15.095
8.3262 | -34.5334 | -15.1053
8.3323 | -34.5338 | -15.1155
8.3384 | -34.5343 | -15.1259
8.3445 | -34.5347 | -15.136
8.3506 | -34.5351 | -15.1463
8.3567 | -34.5356 | -15.1567
8.3628 | -34.536 | -15.1669
8.3689 | -34.5364 | -15.1769
8.375 | -34.5368 | -15.1873
8.3812 | -34.5372 | -15.1974
8.3873 | -34.5376 | -15.2076
8.3934 | -34.5381 | -15.218
8.3995 | -34.5385 | -15.2281
8.4056 | -34.5389 | -15.2383
8.4117 | -34.5393 | -15.2486
8.4178 | -34.5397 | -15.2589
8.4239 | -34.5401 | -15.2689
8.43 | -34.5405 | -15.2791
8.4361 | -34.5409 | -15.2895
8.4422 | -34.5414 | -15.2995
8.4483 | -34.5418 | -15.3096
8.4544 | -34.5422 | -15.3199
8.4605 | -34.5426 | -15.33
8.4666 | -34.5431 | -15.3404
8.4727 | -34.5435 | -15.3507
8.4788 | -34.5439 | -15.3609
8.4849 | -34.5443 | -15.371
8.491 | -34.5447 | -15.3812
8.4971 | -34.5452 | -15.3912
8.5032 | -34.5456 | -15.4013
8.5093 | -34.546 | -15.4113
8.5154 | -34.5464 | -15.4216
8.5215 | -34.5469 | -15.4317
8.5277 | -34.5472 | -15.4419
8.5338 | -34.5477 | -15.452
8.5399 | -34.5481 | -15.4622
8.546 | -34.5485 | -15.4721
8.5521 | -34.549 | -15.4823
8.5582 | -34.5493 | -15.4924
8.5643 | -34.5498 | -15.5023
8.5704 | -34.5502 | -15.5122
8.5765 | -34.5506 | -15.5225
8.5826 | -34.5511 | -15.5323
8.5887 | -34.5515 | -15.5424
8.5948 | -34.5519 | -15.5526
8.6009 | -34.5523 | -15.5625
8.607 | -34.5528 | -15.5728
8.6131 | -34.5532 | -15.583
8.6192 | -34.5536 | -15.5929
8.6253 | -34.5541 | -15.603
8.6314 | -34.5545 | -15.613
8.6375 | -34.5549 | -15.6232
8.6436 | -34.5553 | -15.6333
8.6497 | -34.5557 | -15.6432
8.6558 | -34.5561 | -15.6533
8.6619 | -34.5565 | -15.6634
8.6681 | -34.5569 | -15.6737
8.6742 | -34.5574 | -15.6838
8.6803 | -34.5578 | -15.6939
8.6864 | -34.5582 | -15.7036
8.6925 | -34.5586 | -15.7136
8.6986 | -34.559 | -15.7237
8.7047 | -34.5594 | -15.734
8.7108 | -34.5599 | -15.744
8.7169 | -34.5603 | -15.7541
8.723 | -34.5607 | -15.7639
8.7291 | -34.5611 | -15.7738
8.7352 | -34.5616 | -15.7841
8.7413 | -34.562 | -15.7942
8.7474 | -34.5624 | -15.8041
8.7535 | -34.5628 | -15.8141
8.7596 | -34.5633 | -15.8242
8.7657 | -34.5637 | -15.8344
8.7718 | -34.5641 | -15.8445
8.7779 | -34.5645 | -15.8547
8.784 | -34.5649 | -15.8647
8.7901 | -34.5654 | -15.8747
8.7962 | -34.5658 | -15.8846
8.8023 | -34.5662 | -15.8945
8.8084 | -34.5666 | -15.9046
8.8146 | -34.5671 | -15.9145
8.8207 | -34.5675 | -15.9245
8.8268 | -34.5679 | -15.9342
8.8329 | -34.5683 | -15.9442
8.839 | -34.5688 | -15.9543
8.8451 | -34.5692 | -15.9642
8.8512 | -34.5696 | -15.9741
8.8573 | -34.57 | -15.9841
8.8634 | -34.5705 | -15.9939
8.8695 | -34.5709 | -16.0042
8.8756 | -34.5713 | -16.0142
8.8817 | -34.5717 | -16.024
8.8878 | -34.5721 | -16.0342
8.8939 | -34.5726 | -16.0442
8.9 | -34.573 | -16.0542
8.9061 | -34.5734 | -16.0642
8.9122 | -34.5738 | -16.0741
8.9183 | -34.5743 | -16.0842
8.9244 | -34.5747 | -16.0943
8.9305 | -34.5752 | -16.104
8.9366 | -34.5756 | -16.114
8.9427 | -34.576 | -16.1241
8.9488 | -34.5765 | -16.134
8.955 | -34.5769 | -16.1439
8.9611 | -34.5773 | -16.154
8.9672 | -34.5778 | -16.1639
8.9733 | -34.5782 | -16.1738
8.9794 | -34.5787 | -16.1836
8.9855 | -34.5791 | -16.1936
8.9916 | -34.5795 | -16.2038
8.9977 | -34.58 | -16.2137
9.0038 | -34.5804 | -16.2235
9.0099 | -34.5809 | -16.2334
9.016 | -34.5813 | -16.2435
9.0221 | -34.5817 | -16.2532
9.0282 | -34.5822 | -16.263
9.0343 | -34.5826 | -16.273
9.0404 | -34.5831 | -16.283
9.0465 | -34.5835 | -16.2931
9.0526 | -34.584 | -16.3029
9.0587 | -34.5844 | -16.3129
9.0648 | -34.5849 | -16.3227
9.0709 | -34.5853 | -16.3326
9.077 | -34.5858 | -16.3426
9.0831 | -34.5862 | -16.3527
9.0892 | -34.5866 | -16.3627
9.0953 | -34.5871 | -16.3728
9.1015 | -34.5875 | -16.3826
9.1076 | -34.588 | -16.3923
9.1137 | -34.5884 | -16.4023
9.1198 | -34.5889 | -16.4122
9.1259 | -34.5893 | -16.4221
9.132 | -34.5898 | -16.4321
9.1381 | -34.5902 | -16.4419
9.1442 | -34.5907 | -16.4516
9.1503 | -34.5911 | -16.4616
9.1564 | -34.5916 | -16.4715
9.1625 | -34.592 | -16.4814
9.1686 | -34.5925 | -16.4912
9.1747 | -34.5929 | -16.5012
9.1808 | -34.5934 | -16.511
9.1869 | -34.5938 | -16.5209
9.193 | -34.5943 | -16.5308
9.1991 | -34.5947 | -16.5408
9.2052 | -34.5952 | -16.5507
9.2113 | -34.5956 | -16.5606
9.2174 | -34.5961 | -16.5703
9.2235 | -34.5966 | -16.5802
9.2296 | -34.5971 | -16.59
9.2357 | -34.5975 | -16.6001
9.2419 | -34.598 | -16.61
9.248 | -34.5985 | -16.6198
9.2541 | -34.5989 | -16.6298
9.2602 | -34.5994 | -16.6398
9.2663 | -34.5998 | -16.6494
9.2724 | -34.6003 | -16.6592
9.2785 | -34.6007 | -16.6691
9.2846 | -34.6012 | -16.6787
9.2907 | -34.6017 | -16.6886
9.2968 | -34.6021 | -16.6986
9.3029 | -34.6026 | -16.7086
9.309 | -34.603 | -16.7184
9.3151 | -34.6034 | -16.7282
9.3212 | -34.6039 | -16.7379
9.3273 | -34.6044 | -16.7477
9.3334 | -34.6049 | -16.7574
9.3395 | -34.6053 | -16.7671
9.3456 | -34.6057 | -16.7768
9.3517 | -34.6062 | -16.7867
9.3578 | -34.6066 | -16.7964
9.3639 | -34.6071 | -16.8061
9.37 | -34.6076 | -16.816
9.3761 | -34.608 | -16.826
9.3822 | -34.6085 | -16.8359
9.3884 | -34.609 | -16.8454
9.3945 | -34.6095 | -16.8553
9.4006 | -34.61 | -16.8652
9.4067 | -34.6104 | -16.875
9.4128 | -34.6109 | -16.885
9.4189 | -34.6113 | -16.895
9.425 | -34.6118 | -16.9047
9.4311 | -34.6123 | -16.9146
9.4372 | -34.6128 | -16.9244
9.4433 | -34.6132 | -16.9342
9.4494 | -34.6137 | -16.9438
9.4555 | -34.6142 | -16.9538
9.4616 | -34.6146 | -16.9635
9.4677 | -34.6151 | -16.9733
9.4738 | -34.6156 | -16.9831
9.4799 | -34.6161 | -16.9929
9.486 | -34.6165 | -17.0027
9.4921 | -34.617 | -17.0124
9.4982 | -34.6175 | -17.022
9.5043 | -34.6179 | -17.0317
9.5104 | -34.6184 | -17.0413
9.5165 | -34.6189 | -17.0511
9.5226 | -34.6193 | -17.061
9.5288 | -34.6198 | -17.0706
9.5349 | -34.6203 | -17.0804
9.541 | -34.6207 | -17.09
9.5471 | -34.6212 | -17.0997
9.5532 | -34.6217 | -17.1095
9.5593 | -34.6222 | -17.1194
9.5654 | -34.6226 | -17.129
9.5715 | -34.6231 | -17.1387
9.5776 | -34.6235 | -17.1486
9.5837 | -34.624 | -17.1584
9.5898 | -34.6244 | -17.1684
9.5959 | -34.6249 | -17.1782
9.602 | -34.6254 | -17.1878
9.6081 | -34.6258 | -17.1976
9.6142 | -34.6263 | -17.2073
9.6203 | -34.6268 | -17.217
9.6264 | -34.6272 | -17.2267
9.6325 | -34.6277 | -17.2365
9.6386 | -34.6282 | -17.246
9.6447 | -34.6286 | -17.2558
9.6508 | -34.6291 | -17.2655
9.6569 | -34.6295 | -17.2752
9.663 | -34.63 | -17.285
9.6691 | -34.6305 | -17.2948
9.6753 | -34.631 | -17.3045
9.6814 | -34.6315 | -17.3141
9.6875 | -34.6319 | -17.3237
9.6936 | -34.6324 | -17.3333
9.6997 | -34.6329 | -17.3432
9.7058 | -34.6334 | -17.3528
9.7119 | -34.6339 | -17.3626
9.718 | -34.6343 | -17.3725
9.7241 | -34.6348 | -17.3821
9.7302 | -34.6353 | -17.3919
9.7363 | -34.6358 | -17.4016
9.7424 | -34.6363 | -17.4111
9.7485 | -34.6367 | -17.4209
9.7546 | -34.6372 | -17.4305
9.7607 | -34.6376 | -17.4405
9.7668 | -34.6381 | -17.4504
9.7729 | -34.6386 | -17.4601
9.779 | -34.6391 | -17.4695
9.7851 | -34.6396 | -17.4794
9.7912 | -34.6401 | -17.4891
9.7973 | -34.6406 | -17.4989
9.8034 | -34.6411 | -17.5084
9.8095 | -34.6416 | -17.5183
9.8157 | -34.6421 | -17.528
9.8218 | -34.6426 | -17.5378
9.8279 | -34.6431 | -17.5473
9.834 | -34.6435 | -17.557
9.8401 | -34.644 | -17.5667
9.8462 | -34.6445 | -17.5762
9.8523 | -34.645 | -17.586
9.8584 | -34.6454 | -17.5958
9.8645 | -34.6459 | -17.6054
9.8706 | -34.6463 | -17.6152
9.8767 | -34.6468 | -17.6248
9.8828 | -34.6473 | -17.6344
9.8889 | -34.6477 | -17.6439
9.895 | -34.6482 | -17.6536
9.9011 | -34.6487 | -17.6632
9.9072 | -34.6492 | -17.673
9.9133 | -34.6496 | -17.6827
9.9194 | -34.6501 | -17.6923
9.9255 | -34.6506 | -17.7019
9.9316 | -34.6511 | -17.7116
9.9377 | -34.6516 | -17.7212
9.9438 | -34.6521 | -17.731
9.9499 | -34.6525 | -17.7409
9.956 | -34.653 | -17.7505
9.9622 | -34.6535 | -17.7603
9.9683 | -34.654 | -17.77
9.9744 | -34.6544 | -17.7798
9.9805 | -34.655 | -17.7895
9.9866 | -34.6554 | -17.7991
9.9927 | -34.6559 | -17.8086
9.9988 | -34.6564 | -17.8184
10.0049 | -34.6569 | -17.828
10.011 | -34.6573 | -17.8377
10.0171 | -34.6578 | -17.8473
10.0232 | -34.6583 | -17.857
10.0293 | -34.6589 | -17.8667
10.0354 | -34.6594 | -17.8763
10.0415 | -34.6599 | -17.8859
10.0476 | -34.6603 | -17.8956
10.0537 | -34.6608 | -17.9056
10.0598 | -34.6613 | -17.9153
10.0659 | -34.6617 | -17.925
10.072 | -34.6622 | -17.9344
10.0781 | -34.6627 | -17.944
10.0842 | -34.6632 | -17.9537
10.0903 | -34.6637 | -17.9633
10.0964 | -34.6642 | -17.9726
10.1026 | -34.6646 | -17.9823
10.1087 | -34.6651 | -17.992
10.1148 | -34.6656 | -18.0016
10.1209 | -34.6661 | -18.0112
10.127 | -34.6666 | -18.0206
10.1331 | -34.6671 | -18.0301
10.1392 | -34.6676 | -18.0395
10.1453 | -34.6681 | -18.0492
10.1514 | -34.6686 | -18.0588
10.1575 | -34.669 | -18.0685
10.1636 | -34.6695 | -18.0779
10.1697 | -34.67 | -18.0876
10.1758 | -34.6705 | -18.0971
10.1819 | -34.671 | -18.1066
10.188 | -34.6714 | -18.1162
10.1941 | -34.6719 | -18.1258
10.2002 | -34.6724 | -18.1356
10.2063 | -34.6729 | -18.1451
10.2124 | -34.6734 | -18.1549
10.2185 | -34.6738 | -18.1646
10.2246 | -34.6743 | -18.1741
10.2307 | -34.6748 | -18.1836
10.2368 | -34.6753 | -18.1929
10.2429 | -34.6758 | -18.2026
10.2491 | -34.6762 | -18.2124
10.2552 | -34.6767 | -18.2221
10.2613 | -34.6772 | -18.2317
10.2674 | -34.6777 | -18.2412
10.2735 | -34.6782 | -18.2509
10.2796 | -34.6787 | -18.2603
10.2857 | -34.6792 | -18.2701
10.2918 | -34.6796 | -18.2799
10.2979 | -34.6801 | -18.2895
10.304 | -34.6806 | -18.2989
10.3101 | -34.6811 | -18.3085
10.3162 | -34.6816 | -18.3182
10.3223 | -34.682 | -18.3277
10.3284 | -34.6825 | -18.3374
10.3345 | -34.683 | -18.3467
10.3406 | -34.6835 | -18.3564
10.3467 | -34.6839 | -18.366
10.3528 | -34.6844 | -18.3755
10.3589 | -34.6849 | -18.3851
10.365 | -34.6854 | -18.3947
10.3711 | -34.6859 | -18.4043
10.3772 | -34.6864 | -18.4141
10.3833 | -34.6869 | -18.4236
10.3895 | -34.6873 | -18.4332
10.3956 | -34.6878 | -18.4427
10.4017 | -34.6883 | -18.452
10.4078 | -34.6888 | -18.4616
10.4139 | -34.6893 | -18.4712
10.42 | -34.6898 | -18.4808
10.4261 | -34.6903 | -18.4903
10.4322 | -34.6908 | -18.4999
10.4383 | -34.6912 | -18.5095
10.4444 | -34.6917 | -18.5189
10.4505 | -34.6922 | -18.5285
10.4566 | -34.6927 | -18.5379
10.4627 | -34.6932 | -18.5473
10.4688 | -34.6937 | -18.5568
10.4749 | -34.6942 | -18.5664
10.481 | -34.6947 | -18.5759
10.4871 | -34.6952 | -18.5853
10.4932 | -34.6957 | -18.5947
10.4993 | -34.6962 | -18.6043
10.5054 | -34.6967 | -18.6141
10.5115 | -34.6972 | -18.6239
10.5176 | -34.6977 | -18.6333
10.5237 | -34.6982 | -18.6431
10.5298 | -34.6986 | -18.6525
10.536 | -34.6991 | -18.6622
10.5421 | -34.6996 | -18.6718
10.5482 | -34.7001 | -18.6815
10.5543 | -34.7006 | -18.6909
10.5604 | -34.7011 | -18.7003
10.5665 | -34.7016 | -18.71
10.5726 | -34.7021 | -18.7194
10.5787 | -34.7026 | -18.7288
10.5848 | -34.703 | -18.7381
10.5909 | -34.7035 | -18.7477
10.597 | -34.704 | -18.7573
10.6031 | -34.7044 | -18.7671
10.6092 | -34.7049 | -18.7764
10.6153 | -34.7054 | -18.7864
10.6214 | -34.7059 | -18.7961
10.6275 | -34.7064 | -18.8055
10.6336 | -34.7069 | -18.815
10.6397 | -34.7074 | -18.8248
10.6458 | -34.7079 | -18.8343
10.6519 | -34.7084 | -18.844
10.658 | -34.7089 | -18.8536
10.6641 | -34.7094 | -18.8633
10.6702 | -34.7099 | -18.8728
10.6764 | -34.7104 | -18.8825
10.6825 | -34.7109 | -18.892
10.6886 | -34.7113 | -18.9013
10.6947 | -34.7118 | -18.9109
10.7008 | -34.7123 | -18.9206
10.7069 | -34.7128 | -18.9303
10.713 | -34.7133 | -18.9398
10.7191 | -34.7137 | -18.9493
10.7252 | -34.7142 | -18.9585
10.7313 | -34.7147 | -18.9677
10.7374 | -34.7152 | -18.9772
10.7435 | -34.7157 | -18.9869
10.7496 | -34.7161 | -18.9961
10.7557 | -34.7166 | -19.0058
10.7618 | -34.7171 | -19.0145
10.7679 | -34.7177 | -19.0242
10.774 | -34.7182 | -19.0337
10.7801 | -34.7187 | -19.0432
10.7862 | -34.7192 | -19.0526
10.7923 | -34.7196 | -19.0621
10.7984 | -34.7202 | -19.0722
10.8045 | -34.7207 | -19.0816
10.8106 | -34.7211 | -19.0912
10.8168 | -34.7216 | -19.101
10.8229 | -34.7221 | -19.1108
10.829 | -34.7226 | -19.1199
10.8351 | -34.7232 | -19.1296
10.8412 | -34.7237 | -19.139
10.8473 | -34.7242 | -19.1486
10.8534 | -34.7247 | -19.1583
10.8595 | -34.7252 | -19.1678
10.8656 | -34.7257 | -19.177
10.8717 | -34.7262 | -19.1865
10.8778 | -34.7268 | -19.1961
10.8839 | -34.7273 | -19.2059
10.89 | -34.7278 | -19.2153
10.8961 | -34.7283 | -19.2248
10.9022 | -34.7288 | -19.2345
10.9083 | -34.7293 | -19.2439
10.9144 | -34.7298 | -19.2535
10.9205 | -34.7302 | -19.263
10.9266 | -34.7307 | -19.2723
10.9327 | -34.7312 | -19.2816
10.9388 | -34.7317 | -19.2908
10.9449 | -34.7322 | -19.3003
10.951 | -34.7326 | -19.3097
10.9571 | -34.7331 | -19.3192
10.9633 | -34.7336 | -19.3288
10.9694 | -34.7341 | -19.3386
10.9755 | -34.7346 | -19.3483
10.9816 | -34.735 | -19.3578
10.9877 | -34.7355 | -19.3669
10.9938 | -34.736 | -19.3765
10.9999 | -34.7365 | -19.3861
11.006 | -34.737 | -19.3956
11.0121 | -34.7375 | -19.4051
11.0182 | -34.738 | -19.415
11.0243 | -34.7385 | -19.4243
11.0304 | -34.739 | -19.4339
11.0365 | -34.7395 | -19.443
11.0426 | -34.7401 | -19.4527
11.0487 | -34.7405 | -19.4625
11.0548 | -34.741 | -19.4719
11.0609 | -34.7415 | -19.4814
11.067 | -34.742 | -19.4909
11.0731 | -34.7425 | -19.5004
11.0792 | -34.7429 | -19.51
11.0853 | -34.7434 | -19.5195
11.0914 | -34.7439 | -19.5294
11.0975 | -34.7444 | -19.5385
11.1037 | -34.7449 | -19.5478
11.1098 | -34.7453 | -19.5575
11.1159 | -34.7459 | -19.5667
11.122 | -34.7463 | -19.5763
11.1281 | -34.7468 | -19.5858
11.1342 | -34.7473 | -19.5954
11.1403 | -34.7478 | -19.6047
11.1464 | -34.7483 | -19.6141
11.1525 | -34.7487 | -19.6235
11.1586 | -34.7492 | -19.6328
11.1647 | -34.7497 | -19.6422
11.1708 | -34.7502 | -19.6519
11.1769 | -34.7507 | -19.6615
11.183 | -34.7512 | -19.6708
11.1891 | -34.7517 | -19.6805
11.1952 | -34.7522 | -19.6899
11.2013 | -34.7527 | -19.6994
11.2074 | -34.7532 | -19.7085
11.2135 | -34.7537 | -19.7182
11.2196 | -34.7541 | -19.7279
11.2257 | -34.7547 | -19.7376
11.2318 | -34.7552 | -19.7471
11.2379 | -34.7557 | -19.7568
11.244 | -34.7562 | -19.7664
11.2502 | -34.7567 | -19.7757
11.2563 | -34.7571 | -19.7857
11.2624 | -34.7576 | -19.7953
11.2685 | -34.7582 | -19.8047
11.2746 | -34.7586 | -19.8141
11.2807 | -34.7591 | -19.8237
11.2868 | -34.7597 | -19.8328
11.2929 | -34.7602 | -19.8416
11.299 | -34.7607 | -19.8511
11.3051 | -34.7612 | -19.8609
11.3112 | -34.7616 | -19.8705
11.3173 | -34.7621 | -19.8799
11.3234 | -34.7627 | -19.8897
11.3295 | -34.7632 | -19.8993
11.3356 | -34.7636 | -19.9089
11.3417 | -34.7641 | -19.9185
11.3478 | -34.7645 | -19.9283
11.3539 | -34.765 | -19.9374
11.36 | -34.7655 | -19.9467
11.3661 | -34.766 | -19.956
11.3722 | -34.7665 | -19.9652
11.3783 | -34.767 | -19.9747
11.3844 | -34.7675 | -19.9844
11.3906 | -34.768 | -19.9942
11.3967 | -34.7685 | -20.0038
11.4028 | -34.7689 | -20.0134
11.4089 | -34.7693 | -20.0227
11.415 | -34.7698 | -20.0322
11.4211 | -34.7703 | -20.0415
11.4272 | -34.7708 | -20.0509
11.4333 | -34.7713 | -20.0601
11.4394 | -34.7718 | -20.0696
11.4455 | -34.7723 | -20.0793
11.4516 | -34.7728 | -20.0888
11.4577 | -34.7733 | -20.0983
11.4638 | -34.7739 | -20.1077
11.4699 | -34.7744 | -20.1173
11.476 | -34.7748 | -20.1271
11.4821 | -34.7754 | -20.1368
11.4882 | -34.7759 | -20.1463
11.4943 | -34.7764 | -20.1554
11.5004 | -34.7768 | -20.1643
11.5065 | -34.7772 | -20.1736
11.5126 | -34.7777 | -20.1829
11.5187 | -34.7782 | -20.1922
11.5248 | -34.7787 | -20.2016
11.5309 | -34.7792 | -20.2109
11.5371 | -34.7797 | -20.2202
11.5432 | -34.7802 | -20.2296
11.5493 | -34.7807 | -20.2394
11.5554 | -34.7811 | -20.2487
11.5615 | -34.7816 | -20.2583
11.5676 | -34.7821 | -20.2679
11.5737 | -34.7826 | -20.2774
11.5798 | -34.7832 | -20.2868
11.5859 | -34.7836 | -20.2961
11.592 | -34.7841 | -20.3056
11.5981 | -34.7846 | -20.3151
11.6042 | -34.7851 | -20.3245
11.6103 | -34.7855 | -20.334
11.6164 | -34.7861 | -20.3431
11.6225 | -34.7866 | -20.3524
11.6286 | -34.7871 | -20.3615
11.6347 | -34.7875 | -20.3711
11.6408 | -34.7881 | -20.3806
11.6469 | -34.7885 | -20.39
11.653 | -34.789 | -20.3996
11.6591 | -34.7895 | -20.409
11.6652 | -34.7899 | -20.4183
11.6713 | -34.7904 | -20.4275
11.6775 | -34.791 | -20.4372
11.6836 | -34.7915 | -20.4464
11.6897 | -34.792 | -20.4557
11.6958 | -34.7925 | -20.465
11.7019 | -34.793 | -20.4746
11.708 | -34.7935 | -20.4839
11.7141 | -34.794 | -20.4932
11.7202 | -34.7945 | -20.5027
11.7263 | -34.7949 | -20.5123
11.7324 | -34.7954 | -20.5217
11.7385 | -34.7958 | -20.531
11.7446 | -34.7963 | -20.5401
11.7507 | -34.7968 | -20.5496
11.7568 | -34.7972 | -20.5589
11.7629 | -34.7977 | -20.5681
11.769 | -34.7982 | -20.5773
11.7751 | -34.7987 | -20.5866
11.7812 | -34.7992 | -20.5954
11.7873 | -34.7997 | -20.6049
11.7934 | -34.8002 | -20.6143
11.7995 | -34.8008 | -20.6237
11.8056 | -34.8013 | -20.633
11.8117 | -34.8018 | -20.6423
11.8178 | -34.8023 | -20.6514
11.824 | -34.8028 | -20.6605
11.8301 | -34.8033 | -20.6699
11.8362 | -34.8038 | -20.6793
11.8423 | -34.8043 | -20.6888
11.8484 | -34.8048 | -20.698
11.8545 | -34.8053 | -20.707
11.8606 | -34.8057 | -20.7163
11.8667 | -34.8062 | -20.7259
11.8728 | -34.8068 | -20.7352
11.8789 | -34.8072 | -20.7447
11.885 | -34.8077 | -20.7538
11.8911 | -34.8083 | -20.763
11.8972 | -34.8088 | -20.7722
11.9033 | -34.8092 | -20.7815
11.9094 | -34.8097 | -20.7905
11.9155 | -34.8102 | -20.7997
11.9216 | -34.8107 | -20.8089
11.9277 | -34.8112 | -20.818
11.9338 | -34.8116 | -20.8273
11.9399 | -34.8121 | -20.8366
11.946 | -34.8126 | -20.8458
11.9521 | -34.8131 | -20.8551
11.9582 | -34.8136 | -20.8642
11.9644 | -34.8141 | -20.8734
11.9705 | -34.8146 | -20.8824
11.9766 | -34.8151 | -20.8917
11.9827 | -34.8156 | -20.901
11.9888 | -34.8161 | -20.9103
11.9949 | -34.8166 | -20.9193
12.001 | -34.8171 | -20.9286
12.0071 | -34.8176 | -20.9378
12.0132 | -34.818 | -20.9471
12.0193 | -34.8185 | -20.9565
12.0254 | -34.819 | -20.9655
12.0315 | -34.8195 | -20.9745
12.0376 | -34.82 | -20.9836
12.0437 | -34.8205 | -20.9931
12.0498 | -34.8209 | -21.0025
12.0559 | -34.8215 | -21.0117
12.062 | -34.822 | -21.0207
12.0681 | -34.8225 | -21.0299
12.0742 | -34.823 | -21.0391
12.0803 | -34.8235 | -21.0484
12.0864 | -34.824 | -21.0577
12.0925 | -34.8245 | -21.0669
12.0986 | -34.825 | -21.0759
12.1047 | -34.8255 | -21.0851
12.1109 | -34.826 | -21.0941
12.117 | -34.8265 | -21.1033
12.1231 | -34.827 | -21.1125
12.1292 | -34.8275 | -21.1216
12.1353 | -34.828 | -21.1306
12.1414 | -34.8285 | -21.1399
12.1475 | -34.829 | -21.1492
12.1536 | -34.8295 | -21.1584
12.1597 | -34.8299 | -21.1676
12.1658 | -34.8305 | -21.1765
12.1719 | -34.831 | -21.1855
12.178 | -34.8314 | -21.1947
12.1841 | -34.8319 | -21.2039
12.1902 | -34.8325 | -21.2131
12.1963 | -34.833 | -21.2223
12.2024 | -34.8335 | -21.231
12.2085 | -34.834 | -21.2402
12.2146 | -34.8345 | -21.2494
12.2207 | -34.835 | -21.2584
12.2268 | -34.8355 | -21.2674
12.2329 | -34.836 | -21.2764
12.239 | -34.8365 | -21.2852
12.2451 | -34.837 | -21.2943
12.2513 | -34.8375 | -21.3035
12.2574 | -34.838 | -21.3127
12.2635 | -34.8385 | -21.3219
12.2696 | -34.8389 | -21.3311
12.2757 | -34.8395 | -21.3399
12.2818 | -34.84 | -21.3492
12.2879 | -34.8405 | -21.3582
12.294 | -34.841 | -21.3672
12.3001 | -34.8415 | -21.3761
12.3062 | -34.842 | -21.3852
12.3123 | -34.8425 | -21.394
12.3184 | -34.843 | -21.403
12.3245 | -34.8435 | -21.4118
12.3306 | -34.844 | -21.4208
12.3367 | -34.8445 | -21.4297
12.3428 | -34.845 | -21.4388
12.3489 | -34.8455 | -21.4475
12.355 | -34.846 | -21.4566
12.3611 | -34.8466 | -21.4656
12.3672 | -34.8471 | -21.4745
12.3733 | -34.8476 | -21.4835
12.3794 | -34.8481 | -21.4921
12.3855 | -34.8486 | -21.5011
12.3916 | -34.8491 | -21.5102
12.3978 | -34.8496 | -21.5194
12.4039 | -34.8501 | -21.5285
12.41 | -34.8506 | -21.5375
12.4161 | -34.8511 | -21.546
12.4222 | -34.8516 | -21.555
12.4283 | -34.8521 | -21.5639
12.4344 | -34.8527 | -21.573
12.4405 | -34.8531 | -21.582
12.4466 | -34.8536 | -21.591
12.4527 | -34.8541 | -21.5998
12.4588 | -34.8546 | -21.6086
12.4649 | -34.8552 | -21.6177
12.471 | -34.8557 | -21.6268
12.4771 | -34.8562 | -21.6359
12.4832 | -34.8567 | -21.645
12.4893 | -34.8572 | -21.6535
12.4954 | -34.8577 | -21.6624
12.5015 | -34.8582 | -21.6715
12.5076 | -34.8587 | -21.6807
12.5137 | -34.8592 | -21.6897
12.5198 | -34.8598 | -21.6987
12.5259 | -34.8603 | -21.7078
12.532 | -34.8608 | -21.7167
12.5382 | -34.8613 | -21.7257
12.5443 | -34.8618 | -21.7347
12.5504 | -34.8622 | -21.7436
12.5565 | -34.8628 | -21.7524
12.5626 | -34.8633 | -21.7615
12.5687 | -34.8638 | -21.7705
12.5748 | -34.8644 | -21.7794
12.5809 | -34.8648 | -21.7884
12.587 | -34.8654 | -21.7974
12.5931 | -34.866 | -21.806
12.5992 | -34.8665 | -21.8149
12.6053 | -34.867 | -21.8238
12.6114 | -34.8675 | -21.8326
12.6175 | -34.868 | -21.8416
12.6236 | -34.8685 | -21.8505
12.6297 | -34.869 | -21.8592
12.6358 | -34.8696 | -21.8679
12.6419 | -34.8701 | -21.8769
12.648 | -34.8706 | -21.886
12.6541 | -34.8711 | -21.8948
12.6602 | -34.8716 | -21.9033
12.6663 | -34.8721 | -21.9122
12.6724 | -34.8727 | -21.921
12.6785 | -34.8732 | -21.9299
12.6847 | -34.8737 | -21.9387
12.6908 | -34.8742 | -21.9475
12.6969 | -34.8748 | -21.9564
12.703 | -34.8753 | -21.9654
12.7091 | -34.8758 | -21.9745
12.7152 | -34.8763 | -21.9833
12.7213 | -34.8769 | -21.9922
12.7274 | -34.8774 | -22.0012
12.7335 | -34.8779 | -22.0098
12.7396 | -34.8785 | -22.0185
12.7457 | -34.879 | -22.0274
12.7518 | -34.8795 | -22.0361
12.7579 | -34.8801 | -22.0451
12.764 | -34.8806 | -22.0542
12.7701 | -34.8811 | -22.0629
12.7762 | -34.8816 | -22.0716
12.7823 | -34.8822 | -22.0805
12.7884 | -34.8827 | -22.0893
12.7945 | -34.8833 | -22.0983
12.8006 | -34.8839 | -22.1068
12.8067 | -34.8843 | -22.1156
12.8128 | -34.8848 | -22.1244
12.8189 | -34.8854 | -22.1332
12.8251 | -34.8859 | -22.142
12.8312 | -34.8864 | -22.1508
12.8373 | -34.8869 | -22.1595
12.8434 | -34.8875 | -22.1681
12.8495 | -34.888 | -22.1767
12.8556 | -34.8884 | -22.1854
12.8617 | -34.889 | -22.1942
12.8678 | -34.8895 | -22.203
12.8739 | -34.8901 | -22.2116
12.88 | -34.8906 | -22.2204
12.8861 | -34.8911 | -22.2291
12.8922 | -34.8916 | -22.2379
12.8983 | -34.8922 | -22.2469
12.9044 | -34.8927 | -22.2555
12.9105 | -34.8933 | -22.2641
12.9166 | -34.8938 | -22.2729
12.9227 | -34.8943 | -22.2816
12.9288 | -34.8949 | -22.2903
12.9349 | -34.8954 | -22.2992
12.941 | -34.896 | -22.3079
12.9471 | -34.8966 | -22.3168
12.9532 | -34.8971 | -22.3256
12.9593 | -34.8976 | -22.3345
12.9654 | -34.8982 | -22.3434
12.9716 | -34.8987 | -22.3522
12.9777 | -34.8993 | -22.361
12.9838 | -34.8998 | -22.3697
12.9899 | -34.9003 | -22.3785
12.996 | -34.9008 | -22.3873
13.0021 | -34.9014 | -22.396
13.0082 | -34.9019 | -22.4044
13.0143 | -34.9025 | -22.4129
13.0204 | -34.9031 | -22.4217
13.0265 | -34.9037 | -22.4303
13.0326 | -34.9042 | -22.4389
13.0387 | -34.9048 | -22.4477
13.0448 | -34.9053 | -22.4564
13.0509 | -34.9059 | -22.4651
13.057 | -34.9064 | -22.4738
13.0631 | -34.907 | -22.4826
13.0692 | -34.9076 | -22.4913
13.0753 | -34.9081 | -22.5002
13.0814 | -34.9086 | -22.5091
13.0875 | -34.9092 | -22.5174
13.0936 | -34.9097 | -22.5262
13.0997 | -34.9102 | -22.5347
13.1058 | -34.9108 | -22.5435
13.112 | -34.9114 | -22.5523
13.1181 | -34.9119 | -22.5607
13.1242 | -34.9124 | -22.5694
13.1303 | -34.913 | -22.5781
13.1364 | -34.9135 | -22.587
13.1425 | -34.914 | -22.5955
13.1486 | -34.9146 | -22.6042
13.1547 | -34.9151 | -22.6128
13.1608 | -34.9157 | -22.6216
13.1669 | -34.9162 | -22.6302
13.173 | -34.9168 | -22.6389
13.1791 | -34.9173 | -22.6475
13.1852 | -34.9178 | -22.6564
13.1913 | -34.9184 | -22.6648
13.1974 | -34.9189 | -22.6737
13.2035 | -34.9194 | -22.6822
13.2096 | -34.92 | -22.691
13.2157 | -34.9206 | -22.6999
13.2218 | -34.9211 | -22.7087
13.2279 | -34.9217 | -22.7173
13.234 | -34.9223 | -22.726
13.2401 | -34.9228 | -22.7345
13.2462 | -34.9233 | -22.7431
13.2524 | -34.9238 | -22.752
13.2585 | -34.9244 | -22.7605
13.2646 | -34.9249 | -22.7694
13.2707 | -34.9255 | -22.7779
13.2768 | -34.9261 | -22.7866
13.2829 | -34.9266 | -22.7953
13.289 | -34.9272 | -22.8042
13.2951 | -34.9277 | -22.8123
13.3012 | -34.9283 | -22.8211
13.3073 | -34.9289 | -22.8299
13.3134 | -34.9294 | -22.8387
13.3195 | -34.93 | -22.8473
13.3256 | -34.9305 | -22.8559
13.3317 | -34.9311 | -22.8645
13.3378 | -34.9316 | -22.8733
13.3439 | -34.9321 | -22.882
13.35 | -34.9327 | -22.8909
13.3561 | -34.9332 | -22.8993
13.3622 | -34.9338 | -22.9079
13.3683 | -34.9344 | -22.9168
13.3744 | -34.9349 | -22.9254
13.3805 | -34.9355 | -22.9341
13.3866 | -34.936 | -22.9428
13.3927 | -34.9365 | -22.9514
13.3989 | -34.9371 | -22.9604
13.405 | -34.9376 | -22.969
13.4111 | -34.9382 | -22.9775
13.4172 | -34.9387 | -22.9861
13.4233 | -34.9393 | -22.9951
13.4294 | -34.9398 | -23.0038
13.4355 | -34.9404 | -23.0128
13.4416 | -34.941 | -23.0212
13.4477 | -34.9415 | -23.0297
13.4538 | -34.9421 | -23.0385
13.4599 | -34.9427 | -23.0473
13.466 | -34.9432 | -23.0561
13.4721 | -34.9438 | -23.0646
13.4782 | -34.9443 | -23.0733
13.4843 | -34.9449 | -23.082
13.4904 | -34.9455 | -23.0905
13.4965 | -34.946 | -23.0992
13.5026 | -34.9466 | -23.108
13.5087 | -34.9472 | -23.1166
13.5148 | -34.9477 | -23.1255
13.5209 | -34.9483 | -23.1345
13.527 | -34.9488 | -23.143
13.5331 | -34.9494 | -23.1516
13.5393 | -34.95 | -23.1598
13.5454 | -34.9505 | -23.1683
13.5515 | -34.9511 | -23.1775
13.5576 | -34.9516 | -23.1862
13.5637 | -34.9522 | -23.1949
13.5698 | -34.9527 | -23.2035
13.5759 | -34.9533 | -23.2118
13.582 | -34.9539 | -23.2206
13.5881 | -34.9544 | -23.2292
13.5942 | -34.955 | -23.2378
13.6003 | -34.9555 | -23.2462
13.6064 | -34.9561 | -23.255
13.6125 | -34.9567 | -23.264
13.6186 | -34.9573 | -23.2727
13.6247 | -34.9579 | -23.2814
13.6308 | -34.9584 | -23.2898
13.6369 | -34.959 | -23.2986
13.643 | -34.9596 | -23.3077
13.6491 | -34.9602 | -23.3164
13.6552 | -34.9607 | -23.3252
13.6613 | -34.9612 | -23.3338
13.6674 | -34.9617 | -23.3425
13.6735 | -34.9623 | -23.3513
13.6796 | -34.9629 | -23.3599
13.6858 | -34.9635 | -23.3685
13.6919 | -34.964 | -23.3772
13.698 | -34.9646 | -23.3859
13.7041 | -34.9651 | -23.3949
13.7102 | -34.9657 | -23.4038
13.7163 | -34.9663 | -23.4124
13.7224 | -34.9669 | -23.421
13.7285 | -34.9674 | -23.4296
13.7346 | -34.968 | -23.4382
13.7407 | -34.9686 | -23.447
13.7468 | -34.9691 | -23.4553
13.7529 | -34.9697 | -23.4638
13.759 | -34.9702 | -23.4725
13.7651 | -34.9708 | -23.4815
13.7712 | -34.9713 | -23.4902
13.7773 | -34.9719 | -23.4992
13.7834 | -34.9724 | -23.5077
13.7895 | -34.973 | -23.5164
13.7956 | -34.9735 | -23.5252
13.8017 | -34.9741 | -23.5339
13.8078 | -34.9747 | -23.5425
13.8139 | -34.9752 | -23.551
13.82 | -34.9757 | -23.5596
13.8262 | -34.9763 | -23.5684
13.8323 | -34.9769 | -23.5771
13.8384 | -34.9774 | -23.5858
13.8445 | -34.978 | -23.5946
13.8506 | -34.9786 | -23.6033
13.8567 | -34.9792 | -23.6123
13.8628 | -34.9797 | -23.6209
13.8689 | -34.9803 | -23.6297
13.875 | -34.9809 | -23.6382
13.8811 | -34.9814 | -23.6468
13.8872 | -34.9819 | -23.6555
13.8933 | -34.9825 | -23.664
13.8994 | -34.983 | -23.6727
13.9055 | -34.9836 | -23.6815
13.9116 | -34.9841 | -23.6902
13.9177 | -34.9847 | -23.699
13.9238 | -34.9852 | -23.7076
13.9299 | -34.9858 | -23.7163
13.936 | -34.9863 | -23.725
13.9421 | -34.9868 | -23.7337
13.9482 | -34.9874 | -23.7424
13.9543 | -34.9879 | -23.751
13.9604 | -34.9885 | -23.7597
13.9665 | -34.9891 | -23.7683
13.9727 | -34.9897 | -23.7773
13.9788 | -34.9902 | -23.7861
13.9849 | -34.9908 | -23.7947
13.991 | -34.9914 | -23.8035
13.9971 | -34.9919 | -23.8121
14.0032 | -34.9924 | -23.8208
14.0093 | -34.993 | -23.8294
14.0154 | -34.9936 | -23.838
14.0215 | -34.9941 | -23.8466
14.0276 | -34.9947 | -23.8551
14.0337 | -34.9953 | -23.8637
14.0398 | -34.9959 | -23.8723
14.0459 | -34.9965 | -23.8809
14.052 | -34.997 | -23.8897
14.0581 | -34.9976 | -23.8985
14.0642 | -34.9981 | -23.9072
14.0703 | -34.9987 | -23.9158
14.0764 | -34.9993 | -23.9246
14.0825 | -34.9998 | -23.9332
14.0886 | -35.0004 | -23.9418
14.0947 | -35.001 | -23.9504
14.1008 | -35.0015 | -23.9589
14.1069 | -35.0021 | -23.9674
14.1131 | -35.0026 | -23.976
14.1192 | -35.0032 | -23.9847
14.1253 | -35.0037 | -23.9932
14.1314 | -35.0043 | -24.0017
14.1375 | -35.0049 | -24.0103
14.1436 | -35.0054 | -24.0189
14.1497 | -35.006 | -24.0275
14.1558 | -35.0065 | -24.0361
14.1619 | -35.0071 | -24.0447
14.168 | -35.0076 | -24.0532
14.1741 | -35.0082 | -24.0617
14.1802 | -35.0088 | -24.0704
14.1863 | -35.0093 | -24.0791
14.1924 | -35.0099 | -24.0877
14.1985 | -35.0105 | -24.0964
14.2046 | -35.0111 | -24.1052
14.2107 | -35.0116 | -24.1137
14.2168 | -35.0122 | -24.1224
14.2229 | -35.0127 | -24.131
14.229 | -35.0133 | -24.1399
14.2351 | -35.0138 | -24.1485
14.2412 | -35.0144 | -24.1571
14.2473 | -35.015 | -24.1656
14.2534 | -35.0155 | -24.174
14.2596 | -35.0161 | -24.1825
14.2657 | -35.0166 | -24.1912
14.2718 | -35.0172 | -24.1998
14.2779 | -35.0178 | -24.2086
14.284 | -35.0183 | -24.2172
14.2901 | -35.0189 | -24.2256
14.2962 | -35.0194 | -24.2343
14.3023 | -35.02 | -24.243
14.3084 | -35.0206 | -24.2516
14.3145 | -35.0211 | -24.2601
14.3206 | -35.0217 | -24.269
14.3267 | -35.0222 | -24.2776
14.3328 | -35.0228 | -24.2863
14.3389 | -35.0234 | -24.2949
14.345 | -35.0239 | -24.3035
14.3511 | -35.0244 | -24.3122
14.3572 | -35.025 | -24.321
14.3633 | -35.0255 | -24.3298
14.3694 | -35.0261 | -24.3383
14.3755 | -35.0267 | -24.3471
14.3816 | -35.0273 | -24.3559
14.3877 | -35.0278 | -24.3643
14.3938 | -35.0284 | -24.3731
14.4 | -35.0289 | -24.3817
14.4061 | -35.0294 | -24.3903
14.4122 | -35.03 | -24.399
14.4183 | -35.0305 | -24.4075
14.4244 | -35.0311 | -24.4162
14.4305 | -35.0316 | -24.4248
14.4366 | -35.0322 | -24.4334
14.4427 | -35.0328 | -24.4421
14.4488 | -35.0333 | -24.4508
14.4549 | -35.0339 | -24.4593
14.461 | -35.0345 | -24.4679
14.4671 | -35.035 | -24.4765
14.4732 | -35.0356 | -24.4852
14.4793 | -35.0361 | -24.4938
14.4854 | -35.0367 | -24.5025
14.4915 | -35.0373 | -24.5108
14.4976 | -35.0379 | -24.5193
14.5037 | -35.0384 | -24.5278
14.5098 | -35.039 | -24.5365
14.5159 | -35.0395 | -24.5452
14.522 | -35.0401 | -24.5537
14.5281 | -35.0406 | -24.5621
14.5342 | -35.0412 | -24.571
14.5403 | -35.0418 | -24.5796
14.5465 | -35.0423 | -24.5881
14.5526 | -35.0428 | -24.5969
14.5587 | -35.0434 | -24.6055
14.5648 | -35.044 | -24.6139
14.5709 | -35.0445 | -24.6224
14.577 | -35.0451 | -24.6309
14.5831 | -35.0456 | -24.6395
14.5892 | -35.0462 | -24.6482
14.5953 | -35.0468 | -24.6565
14.6014 | -35.0473 | -24.6652
14.6075 | -35.0479 | -24.6738
14.6136 | -35.0484 | -24.6825
14.6197 | -35.0489 | -24.6911
14.6258 | -35.0494 | -24.6996
14.6319 | -35.05 | -24.708
14.638 | -35.0505 | -24.7167
14.6441 | -35.0511 | -24.7253
14.6502 | -35.0517 | -24.734
14.6563 | -35.0523 | -24.7426
14.6624 | -35.0528 | -24.7514
14.6685 | -35.0534 | -24.7597
14.6746 | -35.0539 | -24.7683
14.6807 | -35.0545 | -24.7767
14.6869 | -35.0551 | -24.7853
14.693 | -35.0556 | -24.7937
14.6991 | -35.0562 | -24.8024
14.7052 | -35.0567 | -24.8109
14.7113 | -35.0573 | -24.8193
14.7174 | -35.0579 | -24.828
14.7235 | -35.0585 | -24.8364
14.7296 | -35.059 | -24.8451
14.7357 | -35.0596 | -24.8534
14.7418 | -35.0601 | -24.8622
14.7479 | -35.0607 | -24.8706
14.754 | -35.0612 | -24.8792
14.7601 | -35.0618 | -24.8878
14.7662 | -35.0624 | -24.8964
14.7723 | -35.0629 | -24.9045
14.7784 | -35.0635 | -24.9132
14.7845 | -35.064 | -24.9219
14.7906 | -35.0645 | -24.9306
14.7967 | -35.0651 | -24.9391
14.8028 | -35.0656 | -24.9477
14.8089 | -35.0662 | -24.9561
14.815 | -35.0668 | -24.9647
14.8211 | -35.0673 | -24.9732
14.8272 | -35.0679 | -24.9817
14.8334 | -35.0684 | -24.9902
14.8395 | -35.069 | -24.9986
14.8456 | -35.0696 | -25.0068
14.8517 | -35.0701 | -25.0154
14.8578 | -35.0707 | -25.024
14.8639 | -35.0712 | -25.0326
14.87 | -35.0718 | -25.0412
14.8761 | -35.0724 | -25.0495
14.8822 | -35.0729 | -25.0581
14.8883 | -35.0734 | -25.0665
14.8944 | -35.0739 | -25.0751
14.9005 | -35.0745 | -25.0837
14.9066 | -35.075 | -25.0922
14.9127 | -35.0755 | -25.1005
14.9188 | -35.0761 | -25.109
14.9249 | -35.0766 | -25.1174
14.931 | -35.0772 | -25.1257
14.9371 | -35.0778 | -25.1339
14.9432 | -35.0783 | -25.1423
14.9493 | -35.0789 | -25.1506
14.9554 | -35.0794 | -25.1591
14.9615 | -35.08 | -25.1674
14.9676 | -35.0805 | -25.1761
14.9738 | -35.0811 | -25.1844
14.9799 | -35.0816 | -25.1929
14.986 | -35.0822 | -25.2012
14.9921 | -35.0827 | -25.2096
14.9982 | -35.0833 | -25.2179
15.0043 | -35.0838 | -25.2263
15.0104 | -35.0844 | -25.2349
15.0165 | -35.0849 | -25.2432
15.0226 | -35.0855 | -25.2513
15.0287 | -35.086 | -25.2599
15.0348 | -35.0866 | -25.2685
15.0409 | -35.0871 | -25.277
15.047 | -35.0877 | -25.2856
15.0531 | -35.0882 | -25.2941
15.0592 | -35.0888 | -25.3025
15.0653 | -35.0894 | -25.3111
15.0714 | -35.09 | -25.3195
15.0775 | -35.0905 | -25.3279
15.0836 | -35.091 | -25.3362
15.0897 | -35.0916 | -25.3443
15.0958 | -35.0922 | -25.3526
15.1019 | -35.0928 | -25.3609
15.108 | -35.0933 | -25.3694
15.1141 | -35.0939 | -25.3778
15.1203 | -35.0944 | -25.3861
15.1264 | -35.095 | -25.3943
15.1325 | -35.0955 | -25.4028
15.1386 | -35.0961 | -25.4111
15.1447 | -35.0966 | -25.4195
15.1508 | -35.0972 | -25.428
15.1569 | -35.0977 | -25.4364
15.163 | -35.0983 | -25.4446
15.1691 | -35.0989 | -25.4531
15.1752 | -35.0994 | -25.4615
15.1813 | -35.1 | -25.4699
15.1874 | -35.1005 | -25.4783
15.1935 | -35.1011 | -25.4865
15.1996 | -35.1017 | -25.4947
15.2057 | -35.1022 | -25.5031
15.2118 | -35.1028 | -25.5114
15.2179 | -35.1034 | -25.5198
15.224 | -35.1039 | -25.5282
15.2301 | -35.1045 | -25.5363
15.2362 | -35.105 | -25.5448
15.2423 | -35.1056 | -25.5533
15.2484 | -35.1062 | -25.5616
15.2545 | -35.1067 | -25.5703
15.2607 | -35.1073 | -25.5788
15.2668 | -35.1078 | -25.5868
15.2729 | -35.1084 | -25.5951
15.279 | -35.109 | -25.6038
15.2851 | -35.1096 | -25.6121
15.2912 | -35.1101 | -25.6205
15.2973 | -35.1107 | -25.6291
15.3034 | -35.1112 | -25.6371
15.3095 | -35.1117 | -25.6454
15.3156 | -35.1122 | -25.6537
15.3217 | -35.1128 | -25.662
15.3278 | -35.1133 | -25.6704
15.3339 | -35.1139 | -25.6784
15.34 | -35.1145 | -25.6868
15.3461 | -35.115 | -25.6948
15.3522 | -35.1156 | -25.7027
15.3583 | -35.1161 | -25.7113
15.3644 | -35.1167 | -25.7195
15.3705 | -35.1173 | -25.7274
15.3766 | -35.1178 | -25.7355
15.3827 | -35.1184 | -25.7437
15.3888 | -35.1189 | -25.7522
15.3949 | -35.1195 | -25.7604
15.401 | -35.12 | -25.7688
15.4072 | -35.1206 | -25.7768
15.4133 | -35.1211 | -25.7851
15.4194 | -35.1217 | -25.7933
15.4255 | -35.1222 | -25.8014
15.4316 | -35.1228 | -25.8099
15.4377 | -35.1234 | -25.8184
15.4438 | -35.1239 | -25.8266
15.4499 | -35.1245 | -25.8349
15.456 | -35.125 | -25.8431
15.4621 | -35.1256 | -25.8513
15.4682 | -35.1261 | -25.8595
15.4743 | -35.1267 | -25.8674
15.4804 | -35.1272 | -25.8758
15.4865 | -35.1277 | -25.8839
15.4926 | -35.1283 | -25.8923
15.4987 | -35.1289 | -25.9004
15.5048 | -35.1295 | -25.9087
15.5109 | -35.13 | -25.9167
15.517 | -35.1306 | -25.9249
15.5231 | -35.1311 | -25.9331
15.5292 | -35.1317 | -25.9415
15.5353 | -35.1322 | -25.95
15.5414 | -35.1328 | -25.9583
15.5476 | -35.1334 | -25.9662
15.5537 | -35.134 | -25.9745
15.5598 | -35.1346 | -25.9828
15.5659 | -35.1351 | -25.9912
15.572 | -35.1357 | -25.9994
15.5781 | -35.1363 | -26.0074
15.5842 | -35.1368 | -26.0153
15.5903 | -35.1374 | -26.0234
15.5964 | -35.138 | -26.0318
15.6025 | -35.1385 | -26.0396
15.6086 | -35.1391 | -26.048
15.6147 | -35.1397 | -26.0565
15.6208 | -35.1403 | -26.0644
15.6269 | -35.1408 | -26.0727
15.633 | -35.1414 | -26.0807
15.6391 | -35.142 | -26.089
15.6452 | -35.1425 | -26.0972
15.6513 | -35.1431 | -26.1051
15.6574 | -35.1436 | -26.1136
15.6635 | -35.1442 | -26.1216
15.6696 | -35.1447 | -26.13
15.6757 | -35.1453 | -26.1381
15.6818 | -35.1458 | -26.1465
15.688 | -35.1464 | -26.1544
15.6941 | -35.1469 | -26.1626
15.7002 | -35.1475 | -26.1706
15.7063 | -35.148 | -26.1786
15.7124 | -35.1486 | -26.1869
15.7185 | -35.1491 | -26.1951
15.7246 | -35.1497 | -26.2031
15.7307 | -35.1503 | -26.2113
15.7368 | -35.1509 | -26.2194
15.7429 | -35.1514 | -26.2277
15.749 | -35.152 | -26.2358
15.7551 | -35.1526 | -26.244
15.7612 | -35.1532 | -26.2518
15.7673 | -35.1538 | -26.2598
15.7734 | -35.1543 | -26.2683
15.7795 | -35.1548 | -26.2764
15.7856 | -35.1554 | -26.2844
15.7917 | -35.156 | -26.2923
15.7978 | -35.1565 | -26.3007
15.8039 | -35.1571 | -26.3088
15.81 | -35.1577 | -26.3169
15.8161 | -35.1582 | -26.3249
15.8222 | -35.1588 | -26.3328
15.8283 | -35.1594 | -26.3408
15.8345 | -35.1599 | -26.349
15.8406 | -35.1605 | -26.3572
15.8467 | -35.1611 | -26.3653
15.8528 | -35.1617 | -26.3734
15.8589 | -35.1623 | -26.3818
15.865 | -35.1628 | -26.3895
15.8711 | -35.1634 | -26.3976
15.8772 | -35.1639 | -26.4055
15.8833 | -35.1644 | -26.4136
15.8894 | -35.165 | -26.4219
15.8955 | -35.1655 | -26.4301
15.9016 | -35.1661 | -26.4378
15.9077 | -35.1666 | -26.4458
15.9138 | -35.1672 | -26.454
15.9199 | -35.1677 | -26.462
15.926 | -35.1682 | -26.4703
15.9321 | -35.1688 | -26.4781
15.9382 | -35.1693 | -26.4862
15.9443 | -35.1699 | -26.4942
15.9504 | -35.1705 | -26.5022
15.9565 | -35.171 | -26.5101
15.9626 | -35.1715 | -26.5182
15.9687 | -35.1721 | -26.526
15.9749 | -35.1727 | -26.5338
15.981 | -35.1732 | -26.5418
15.9871 | -35.1738 | -26.5501
15.9932 | -35.1743 | -26.5578
15.9993 | -35.1749 | -26.5658
16.0054 | -35.1755 | -26.5734
16.0115 | -35.1761 | -26.5812
16.0176 | -35.1766 | -26.5894
16.0237 | -35.1772 | -26.5975
16.0298 | -35.1777 | -26.6054
16.0359 | -35.1782 | -26.6139
16.042 | -35.1788 | -26.6217
16.0481 | -35.1793 | -26.63
16.0542 | -35.1799 | -26.638
16.0603 | -35.1805 | -26.6459
16.0664 | -35.181 | -26.654
16.0725 | -35.1815 | -26.6618
16.0786 | -35.1821 | -26.6696
16.0847 | -35.1827 | -26.6778
16.0908 | -35.1832 | -26.6856
16.0969 | -35.1837 | -26.6937
16.103 | -35.1843 | -26.7018
16.1091 | -35.1849 | -26.7095
16.1152 | -35.1854 | -26.7174
16.1214 | -35.186 | -26.7251
16.1275 | -35.1866 | -26.7331
16.1336 | -35.1872 | -26.7414
16.1397 | -35.1877 | -26.7491
16.1458 | -35.1883 | -26.757
16.1519 | -35.1888 | -26.7654
16.158 | -35.1894 | -26.7735
16.1641 | -35.1899 | -26.7815
16.1702 | -35.1905 | -26.7896
16.1763 | -35.191 | -26.7975
16.1824 | -35.1916 | -26.8053
16.1885 | -35.1921 | -26.813
16.1946 | -35.1927 | -26.821
16.2007 | -35.1933 | -26.8288
16.2068 | -35.1939 | -26.8367
16.2129 | -35.1945 | -26.8448
16.219 | -35.195 | -26.8524
16.2251 | -35.1956 | -26.8602
16.2312 | -35.1961 | -26.8683
16.2373 | -35.1967 | -26.8762
16.2434 | -35.1972 | -26.8844
16.2495 | -35.1978 | -26.8921
16.2556 | -35.1984 | -26.9001
16.2618 | -35.1989 | -26.9078
16.2679 | -35.1994 | -26.9157
16.274 | -35.1999 | -26.9235
16.2801 | -35.2005 | -26.9313
16.2862 | -35.201 | -26.939
16.2923 | -35.2016 | -26.947
16.2984 | -35.2021 | -26.9547
16.3045 | -35.2026 | -26.9627
16.3106 | -35.2032 | -26.9705
16.3167 | -35.2037 | -26.9783
16.3228 | -35.2042 | -26.986
16.3289 | -35.2048 | -26.994
16.335 | -35.2054 | -27.0022
16.3411 | -35.206 | -27.0103
16.3472 | -35.2065 | -27.0185
16.3533 | -35.2071 | -27.0266
16.3594 | -35.2077 | -27.0343
16.3655 | -35.2082 | -27.0421
16.3716 | -35.2088 | -27.0502
16.3777 | -35.2093 | -27.0584
16.3838 | -35.2099 | -27.0664
16.3899 | -35.2105 | -27.0742
16.396 | -35.2111 | -27.082
16.4021 | -35.2116 | -27.0898
16.4083 | -35.2122 | -27.0977
16.4144 | -35.2127 | -27.1053
16.4205 | -35.2133 | -27.1133
16.4266 | -35.2139 | -27.1207
16.4327 | -35.2145 | -27.1283
16.4388 | -35.215 | -27.136
16.4449 | -35.2155 | -27.1435
16.451 | -35.2161 | -27.1515
16.4571 | -35.2166 | -27.1594
16.4632 | -35.2171 | -27.1669
16.4693 | -35.2177 | -27.1749
16.4754 | -35.2182 | -27.1827
16.4815 | -35.2188 | -27.1904
16.4876 | -35.2193 | -27.1982
16.4937 | -35.2199 | -27.2059
16.4998 | -35.2205 | -27.2135
16.5059 | -35.221 | -27.2211
16.512 | -35.2216 | -27.2293
16.5181 | -35.2221 | -27.2373
16.5242 | -35.2226 | -27.2455
16.5303 | -35.2232 | -27.2531
16.5364 | -35.2238 | -27.2607
16.5425 | -35.2243 | -27.2685
16.5487 | -35.2248 | -27.2763
16.5548 | -35.2254 | -27.2842
16.5609 | -35.226 | -27.2919
16.567 | -35.2265 | -27.2992
16.5731 | -35.2271 | -27.3072
16.5792 | -35.2276 | -27.315
16.5853 | -35.2282 | -27.3228
16.5914 | -35.2287 | -27.3305
16.5975 | -35.2293 | -27.3383
16.6036 | -35.2299 | -27.3461
16.6097 | -35.2304 | -27.3542
16.6158 | -35.2309 | -27.3621
16.6219 | -35.2315 | -27.3701
16.628 | -35.232 | -27.3778
16.6341 | -35.2326 | -27.3856
16.6402 | -35.2331 | -27.3931
16.6463 | -35.2336 | -27.4009
16.6524 | -35.2342 | -27.4085
16.6585 | -35.2347 | -27.4161
16.6646 | -35.2353 | -27.4238
16.6707 | -35.2359 | -27.4314
16.6768 | -35.2364 | -27.4391
16.6829 | -35.237 | -27.4469
16.689 | -35.2376 | -27.4549
16.6952 | -35.2381 | -27.4624
16.7013 | -35.2386 | -27.4702
16.7074 | -35.2392 | -27.4777
16.7135 | -35.2398 | -27.4854
16.7196 | -35.2403 | -27.4933
16.7257 | -35.2409 | -27.5011
16.7318 | -35.2415 | -27.509
16.7379 | -35.242 | -27.5168
16.744 | -35.2425 | -27.5243
16.7501 | -35.2431 | -27.5321
16.7562 | -35.2436 | -27.5398
16.7623 | -35.2441 | -27.5474
16.7684 | -35.2446 | -27.5551
16.7745 | -35.2452 | -27.563
16.7806 | -35.2458 | -27.5705
16.7867 | -35.2463 | -27.5783
16.7928 | -35.2469 | -27.586
16.7989 | -35.2474 | -27.5938
16.805 | -35.2479 | -27.6017
16.8111 | -35.2485 | -27.6094
16.8172 | -35.249 | -27.6172
16.8233 | -35.2496 | -27.6248
16.8294 | -35.2501 | -27.6324
16.8356 | -35.2507 | -27.6403
16.8417 | -35.2513 | -27.6482
16.8478 | -35.2518 | -27.6557
16.8539 | -35.2523 | -27.6632
16.86 | -35.2529 | -27.671
16.8661 | -35.2534 | -27.6786
16.8722 | -35.254 | -27.6863
16.8783 | -35.2545 | -27.694
16.8844 | -35.2551 | -27.7017
16.8905 | -35.2556 | -27.7095
16.8966 | -35.2562 | -27.7172
16.9027 | -35.2567 | -27.7249
16.9088 | -35.2573 | -27.7327
16.9149 | -35.2578 | -27.7404
16.921 | -35.2584 | -27.7481
16.9271 | -35.259 | -27.7561
16.9332 | -35.2595 | -27.7639
16.9393 | -35.2601 | -27.7715
16.9454 | -35.2606 | -27.7791
16.9515 | -35.2612 | -27.7869
16.9576 | -35.2617 | -27.7944
16.9637 | -35.2623 | -27.8021
16.9698 | -35.2629 | -27.8097
16.9759 | -35.2634 | -27.8174
16.9821 | -35.264 | -27.8253
16.9882 | -35.2646 | -27.8329
16.9943 | -35.2651 | -27.8406
17.0004 | -35.2656 | -27.8483
17.0065 | -35.2662 | -27.856
17.0126 | -35.2667 | -27.8636
17.0187 | -35.2673 | -27.8714
17.0248 | -35.2679 | -27.879
17.0309 | -35.2684 | -27.8865
17.037 | -35.269 | -27.8941
17.0431 | -35.2695 | -27.9018
17.0492 | -35.2701 | -27.9094
17.0553 | -35.2706 | -27.917
17.0614 | -35.2712 | -27.9245
17.0675 | -35.2717 | -27.9322
17.0736 | -35.2723 | -27.9399
17.0797 | -35.2728 | -27.9474
17.0858 | -35.2733 | -27.9554
17.0919 | -35.2739 | -27.9632
17.098 | -35.2744 | -27.9707
17.1041 | -35.275 | -27.9782
17.1102 | -35.2756 | -27.9859
17.1163 | -35.2761 | -27.9933
17.1225 | -35.2767 | -28.0008
17.1286 | -35.2773 | -28.0081
17.1347 | -35.2778 | -28.0157
17.1408 | -35.2784 | -28.0236
17.1469 | -35.2789 | -28.0311
17.153 | -35.2794 | -28.0387
17.1591 | -35.28 | -28.0463
17.1652 | -35.2805 | -28.0537
17.1713 | -35.281 | -28.0614
17.1774 | -35.2816 | -28.069
17.1835 | -35.2822 | -28.0767
17.1896 | -35.2827 | -28.0844
17.1957 | -35.2832 | -28.092
17.2018 | -35.2838 | -28.0994
17.2079 | -35.2844 | -28.107
17.214 | -35.2849 | -28.1148
17.2201 | -35.2855 | -28.1224
17.2262 | -35.286 | -28.13
17.2323 | -35.2866 | -28.1377
17.2384 | -35.2871 | -28.1451
17.2445 | -35.2877 | -28.1528
17.2506 | -35.2882 | -28.1605
17.2567 | -35.2888 | -28.1682
17.2628 | -35.2893 | -28.1758
17.269 | -35.2899 | -28.1834
17.2751 | -35.2904 | -28.191
17.2812 | -35.291 | -28.1987
17.2873 | -35.2915 | -28.2065
17.2934 | -35.2921 | -28.2141
17.2995 | -35.2926 | -28.2218
17.3056 | -35.2932 | -28.2293
17.3117 | -35.2937 | -28.237
17.3178 | -35.2943 | -28.2446
17.3239 | -35.2948 | -28.2522
17.33 | -35.2953 | -28.2598
17.3361 | -35.2959 | -28.2676
17.3422 | -35.2964 | -28.2752
17.3483 | -35.297 | -28.2828
17.3544 | -35.2975 | -28.2905
17.3605 | -35.2981 | -28.2979
17.3666 | -35.2986 | -28.3056
17.3727 | -35.2992 | -28.3134
17.3788 | -35.2997 | -28.3208
17.3849 | -35.3003 | -28.3283
17.391 | -35.3008 | -28.3359
17.3971 | -35.3013 | -28.3435
17.4032 | -35.3019 | -28.3511
17.4094 | -35.3025 | -28.3586
17.4155 | -35.303 | -28.3661
17.4216 | -35.3035 | -28.3736
17.4277 | -35.3041 | -28.3812
17.4338 | -35.3047 | -28.3888
17.4399 | -35.3052 | -28.3966
17.446 | -35.3058 | -28.4039
17.4521 | -35.3064 | -28.4115
17.4582 | -35.3069 | -28.4189
17.4643 | -35.3074 | -28.4266
17.4704 | -35.308 | -28.4341
17.4765 | -35.3085 | -28.442
17.4826 | -35.3091 | -28.4493
17.4887 | -35.3096 | -28.457
17.4948 | -35.3102 | -28.4644
17.5009 | -35.3107 | -28.472
17.507 | -35.3113 | -28.4796
17.5131 | -35.3119 | -28.4874
17.5192 | -35.3124 | -28.4948
17.5253 | -35.3129 | -28.5024
17.5314 | -35.3134 | -28.51
17.5375 | -35.314 | -28.5177
17.5436 | -35.3145 | -28.5254
17.5497 | -35.315 | -28.533
17.5559 | -35.3156 | -28.5403
17.562 | -35.3161 | -28.5479
17.5681 | -35.3167 | -28.5555
17.5742 | -35.3172 | -28.5632
17.5803 | -35.3178 | -28.5708
17.5864 | -35.3184 | -28.5783
17.5925 | -35.3189 | -28.5858
17.5986 | -35.3195 | -28.5933
17.6047 | -35.32 | -28.6009
17.6108 | -35.3206 | -28.6083
17.6169 | -35.3211 | -28.6158
17.623 | -35.3217 | -28.6232
17.6291 | -35.3222 | -28.6307
17.6352 | -35.3228 | -28.6383
17.6413 | -35.3233 | -28.646
17.6474 | -35.3239 | -28.6536
17.6535 | -35.3244 | -28.6611
17.6596 | -35.325 | -28.6685
17.6657 | -35.3255 | -28.6762
17.6718 | -35.3261 | -28.6838
17.6779 | -35.3266 | -28.6913
17.684 | -35.3272 | -28.6988
17.6901 | -35.3278 | -28.7064
17.6963 | -35.3283 | -28.7137
17.7024 | -35.3288 | -28.7213
17.7085 | -35.3294 | -28.7288
17.7146 | -35.33 | -28.7362
17.7207 | -35.3305 | -28.7435
17.7268 | -35.3311 | -28.7509
17.7329 | -35.3316 | -28.7582
17.739 | -35.3322 | -28.7658
17.7451 | -35.3328 | -28.7733
17.7512 | -35.3333 | -28.781
17.7573 | -35.3339 | -28.7885
17.7634 | -35.3345 | -28.796
17.7695 | -35.335 | -28.8033
17.7756 | -35.3356 | -28.8108
17.7817 | -35.3361 | -28.8181
17.7878 | -35.3367 | -28.8257
17.7939 | -35.3372 | -28.8331
17.8 | -35.3378 | -28.8405
17.8061 | -35.3383 | -28.848
17.8122 | -35.3389 | -28.8554
17.8183 | -35.3394 | -28.863
17.8244 | -35.34 | -28.8704
17.8305 | -35.3405 | -28.878
17.8366 | -35.3411 | -28.8851
17.8428 | -35.3417 | -28.8926
17.8489 | -35.3422 | -28.9
17.855 | -35.3428 | -28.9074
17.8611 | -35.3433 | -28.9148
17.8672 | -35.3439 | -28.9221
17.8733 | -35.3444 | -28.9297
17.8794 | -35.345 | -28.9371
17.8855 | -35.3455 | -28.9447
17.8916 | -35.346 | -28.9523
17.8977 | -35.3466 | -28.9597
17.9038 | -35.3471 | -28.9668
17.9099 | -35.3477 | -28.9744
17.916 | -35.3483 | -28.982
17.9221 | -35.3488 | -28.9894
17.9282 | -35.3493 | -28.9968
17.9343 | -35.3499 | -29.0044
17.9404 | -35.3505 | -29.0117
17.9465 | -35.351 | -29.0192
17.9526 | -35.3516 | -29.0267
17.9587 | -35.3521 | -29.0341
17.9648 | -35.3527 | -29.0414
17.9709 | -35.3532 | -29.0489
17.977 | -35.3537 | -29.0562
17.9832 | -35.3543 | -29.0638
17.9893 | -35.3548 | -29.0714
17.9954 | -35.3554 | -29.0789
18.0015 | -35.3559 | -29.0864
18.0076 | -35.3565 | -29.0937
18.0137 | -35.357 | -29.1012
18.0198 | -35.3575 | -29.1086
18.0259 | -35.3581 | -29.1161
18.032 | -35.3586 | -29.1235
18.0381 | -35.3591 | -29.131
18.0442 | -35.3597 | -29.1384
18.0503 | -35.3602 | -29.146
18.0564 | -35.3608 | -29.1533
18.0625 | -35.3613 | -29.1608
18.0686 | -35.3619 | -29.1683
18.0747 | -35.3624 | -29.1757
18.0808 | -35.3629 | -29.183
18.0869 | -35.3635 | -29.1905
18.093 | -35.3641 | -29.198
18.0991 | -35.3646 | -29.2054
18.1052 | -35.3651 | -29.2131
18.1113 | -35.3657 | -29.2206
18.1174 | -35.3662 | -29.2279
18.1236 | -35.3668 | -29.2355
18.1297 | -35.3673 | -29.2429
18.1358 | -35.3679 | -29.2502
18.1419 | -35.3685 | -29.2577
18.148 | -35.369 | -29.2651
18.1541 | -35.3695 | -29.2722
18.1602 | -35.3701 | -29.2797
18.1663 | -35.3707 | -29.2869
18.1724 | -35.3712 | -29.2945
18.1785 | -35.3717 | -29.3019
18.1846 | -35.3723 | -29.3093
18.1907 | -35.3728 | -29.3168
18.1968 | -35.3733 | -29.3244
18.2029 | -35.3739 | -29.3319
18.209 | -35.3744 | -29.3394
18.2151 | -35.375 | -29.3469
18.2212 | -35.3756 | -29.3541
18.2273 | -35.3761 | -29.3616
18.2334 | -35.3767 | -29.3691
18.2395 | -35.3772 | -29.3764
18.2456 | -35.3778 | -29.3839
18.2517 | -35.3783 | -29.3917
18.2578 | -35.3788 | -29.399
18.2639 | -35.3794 | -29.4067
18.2701 | -35.38 | -29.4142
18.2762 | -35.3805 | -29.4216
18.2823 | -35.3811 | -29.4292
18.2884 | -35.3816 | -29.4365
18.2945 | -35.3822 | -29.4437
18.3006 | -35.3827 | -29.4514
18.3067 | -35.3833 | -29.4589
18.3128 | -35.3839 | -29.4665
18.3189 | -35.3844 | -29.474
18.325 | -35.3849 | -29.4813
18.3311 | -35.3855 | -29.4885
18.3372 | -35.386 | -29.4959
18.3433 | -35.3866 | -29.5034
18.3494 | -35.3871 | -29.5108
18.3555 | -35.3877 | -29.5182
18.3616 | -35.3883 | -29.5254
18.3677 | -35.3888 | -29.5328
18.3738 | -35.3894 | -29.5402
18.3799 | -35.3899 | -29.5476
18.386 | -35.3905 | -29.555
18.3921 | -35.391 | -29.5624
18.3982 | -35.3915 | -29.5694
18.4043 | -35.3921 | -29.5768
18.4105 | -35.3926 | -29.5841
18.4166 | -35.3931 | -29.5913
18.4227 | -35.3937 | -29.5986
18.4288 | -35.3943 | -29.6061
18.4349 | -35.3949 | -29.6133
18.441 | -35.3954 | -29.6207
18.4471 | -35.396 | -29.628
18.4532 | -35.3965 | -29.6354
18.4593 | -35.3971 | -29.6427
18.4654 | -35.3976 | -29.6497
18.4715 | -35.3982 | -29.6572
18.4776 | -35.3987 | -29.6648
18.4837 | -35.3993 | -29.6723
18.4898 | -35.3998 | -29.6798
18.4959 | -35.4003 | -29.6873
18.502 | -35.4009 | -29.6945
18.5081 | -35.4015 | -29.7019
18.5142 | -35.402 | -29.7095
18.5203 | -35.4025 | -29.7169
18.5264 | -35.4031 | -29.7244
18.5325 | -35.4037 | -29.7318
18.5386 | -35.4042 | -29.7393
18.5447 | -35.4048 | -29.7465
18.5508 | -35.4053 | -29.7539
18.557 | -35.4059 | -29.7615
18.5631 | -35.4064 | -29.7687
18.5692 | -35.4069 | -29.7761
18.5753 | -35.4075 | -29.7835
18.5814 | -35.408 | -29.7909
18.5875 | -35.4086 | -29.7982
18.5936 | -35.4092 | -29.8056
18.5997 | -35.4097 | -29.8131
18.6058 | -35.4103 | -29.8204
18.6119 | -35.4108 | -29.8279
18.618 | -35.4114 | -29.8354
18.6241 | -35.4119 | -29.8427
18.6302 | -35.4125 | -29.8503
18.6363 | -35.413 | -29.8575
18.6424 | -35.4136 | -29.8647
18.6485 | -35.4142 | -29.8722
18.6546 | -35.4148 | -29.8796
18.6607 | -35.4153 | -29.8871
18.6668 | -35.4159 | -29.8942
18.6729 | -35.4164 | -29.9018
18.679 | -35.417 | -29.9089
18.6851 | -35.4175 | -29.9162
18.6912 | -35.4181 | -29.9234
18.6974 | -35.4187 | -29.9307
18.7035 | -35.4192 | -29.9379
18.7096 | -35.4198 | -29.9454
18.7157 | -35.4203 | -29.9526
18.7218 | -35.4209 | -29.96
18.7279 | -35.4215 | -29.9673
18.734 | -35.422 | -29.9746
18.7401 | -35.4226 | -29.9817
18.7462 | -35.4232 | -29.989
18.7523 | -35.4237 | -29.9965
18.7584 | -35.4244 | -30.004
18.7645 | -35.4249 | -30.0115
18.7706 | -35.4255 | -30.0187
18.7767 | -35.426 | -30.0257
18.7828 | -35.4266 | -30.0331
18.7889 | -35.4271 | -30.0403
18.795 | -35.4277 | -30.0479
18.8011 | -35.4283 | -30.0552
18.8072 | -35.4288 | -30.0624
18.8133 | -35.4294 | -30.0699
18.8194 | -35.4299 | -30.077
18.8255 | -35.4305 | -30.0844
18.8316 | -35.431 | -30.0918
18.8377 | -35.4316 | -30.0991
18.8439 | -35.4321 | -30.1065
18.85 | -35.4327 | -30.1137
18.8561 | -35.4333 | -30.1208
18.8622 | -35.4338 | -30.1283
18.8683 | -35.4344 | -30.1357
18.8744 | -35.4349 | -30.1431
18.8805 | -35.4355 | -30.1503
18.8866 | -35.4361 | -30.1579
18.8927 | -35.4366 | -30.1651
18.8988 | -35.4372 | -30.1725
18.9049 | -35.4377 | -30.1798
18.911 | -35.4383 | -30.187
18.9171 | -35.4388 | -30.1942
18.9232 | -35.4394 | -30.2013
18.9293 | -35.44 | -30.2088
18.9354 | -35.4405 | -30.2161
18.9415 | -35.4411 | -30.2234
18.9476 | -35.4416 | -30.2305
18.9537 | -35.4422 | -30.2379
18.9598 | -35.4428 | -30.2451
18.9659 | -35.4434 | -30.2524
18.972 | -35.4439 | -30.2597
18.9781 | -35.4445 | -30.267
18.9843 | -35.445 | -30.2744
18.9904 | -35.4456 | -30.2816
18.9965 | -35.4461 | -30.2887
19.0026 | -35.4467 | -30.2959
19.0087 | -35.4472 | -30.3031
19.0148 | -35.4478 | -30.3103
19.0209 | -35.4484 | -30.3175
19.027 | -35.4489 | -30.325
19.0331 | -35.4495 | -30.3321
19.0392 | -35.4501 | -30.3394
19.0453 | -35.4506 | -30.3466
19.0514 | -35.4512 | -30.3537
19.0575 | -35.4517 | -30.3611
19.0636 | -35.4523 | -30.3683
19.0697 | -35.4529 | -30.3756
19.0758 | -35.4534 | -30.3826
19.0819 | -35.454 | -30.39
19.088 | -35.4546 | -30.3968
19.0941 | -35.4551 | -30.4042
19.1002 | -35.4557 | -30.4113
19.1063 | -35.4562 | -30.4187
19.1124 | -35.4568 | -30.4261
19.1185 | -35.4573 | -30.4333
19.1246 | -35.4579 | -30.4403
19.1308 | -35.4585 | -30.4477
19.1369 | -35.459 | -30.4549
19.143 | -35.4596 | -30.4623
19.1491 | -35.4602 | -30.4695
19.1552 | -35.4606 | -30.477
19.1613 | -35.4612 | -30.4842
19.1674 | -35.4617 | -30.4917
19.1735 | -35.4623 | -30.4989
19.1796 | -35.4629 | -30.5063
19.1857 | -35.4634 | -30.5134
19.1918 | -35.464 | -30.5207
19.1979 | -35.4646 | -30.5279
19.204 | -35.4651 | -30.5351
19.2101 | -35.4657 | -30.5423
19.2162 | -35.4662 | -30.5494
19.2223 | -35.4668 | -30.5567
19.2284 | -35.4673 | -30.5639
19.2345 | -35.4679 | -30.5715
19.2406 | -35.4684 | -30.5787
19.2467 | -35.4689 | -30.5861
19.2528 | -35.4695 | -30.5933
19.2589 | -35.4701 | -30.6004
19.265 | -35.4706 | -30.6078
19.2712 | -35.4712 | -30.6151
19.2773 | -35.4718 | -30.6222
19.2834 | -35.4724 | -30.6295
19.2895 | -35.4729 | -30.6368
19.2956 | -35.4734 | -30.6441
19.3017 | -35.474 | -30.6513
19.3078 | -35.4745 | -30.6586
19.3139 | -35.475 | -30.6657
19.32 | -35.4756 | -30.6731
19.3261 | -35.4762 | -30.6805
19.3322 | -35.4767 | -30.6876
19.3383 | -35.4773 | -30.6947
19.3444 | -35.4778 | -30.7017
19.3505 | -35.4784 | -30.7088
19.3566 | -35.4789 | -30.7159
19.3627 | -35.4794 | -30.723
19.3688 | -35.48 | -30.7303
19.3749 | -35.4806 | -30.7375
19.381 | -35.4812 | -30.7444
19.3871 | -35.4817 | -30.7513
19.3932 | -35.4823 | -30.7584
19.3993 | -35.4828 | -30.7657
19.4054 | -35.4834 | -30.7729
19.4115 | -35.484 | -30.7803
19.4177 | -35.4845 | -30.7874
19.4238 | -35.4851 | -30.7946
19.4299 | -35.4856 | -30.8017
19.436 | -35.4862 | -30.8089
19.4421 | -35.4867 | -30.816
19.4482 | -35.4873 | -30.8233
19.4543 | -35.4878 | -30.8302
19.4604 | -35.4884 | -30.8375
19.4665 | -35.4889 | -30.8445
19.4726 | -35.4895 | -30.8514
19.4787 | -35.4901 | -30.8584
19.4848 | -35.4906 | -30.8657
19.4909 | -35.4912 | -30.8726
19.497 | -35.4917 | -30.88
19.5031 | -35.4923 | -30.8871
19.5092 | -35.4928 | -30.8943
19.5153 | -35.4934 | -30.9016
19.5214 | -35.494 | -30.9085
19.5275 | -35.4945 | -30.9155
19.5336 | -35.4951 | -30.9229
19.5397 | -35.4957 | -30.93
19.5458 | -35.4963 | -30.9373
19.5519 | -35.4968 | -30.9445
19.5581 | -35.4974 | -30.9515
19.5642 | -35.4979 | -30.9584
19.5703 | -35.4985 | -30.9654
19.5764 | -35.4991 | -30.9724
19.5825 | -35.4997 | -30.9796
19.5886 | -35.5003 | -30.9868
19.5947 | -35.5008 | -30.9938
19.6008 | -35.5014 | -31.001
19.6069 | -35.5019 | -31.0082
19.613 | -35.5025 | -31.0154
19.6191 | -35.503 | -31.0225
19.6252 | -35.5036 | -31.0298
19.6313 | -35.5041 | -31.0368
19.6374 | -35.5047 | -31.0441
19.6435 | -35.5052 | -31.0515
19.6496 | -35.5058 | -31.0585
19.6557 | -35.5064 | -31.0656
19.6618 | -35.5069 | -31.0724
19.6679 | -35.5075 | -31.0796
19.674 | -35.508 | -31.0865
19.6801 | -35.5086 | -31.0936
19.6862 | -35.5091 | -31.1006
19.6923 | -35.5097 | -31.1075
19.6984 | -35.5102 | -31.1145
19.7046 | -35.5108 | -31.1217
19.7107 | -35.5114 | -31.1287
19.7168 | -35.5119 | -31.1361
19.7229 | -35.5124 | -31.1431
19.729 | -35.513 | -31.1501
19.7351 | -35.5136 | -31.157
19.7412 | -35.5141 | -31.1643
19.7473 | -35.5147 | -31.1715
19.7534 | -35.5153 | -31.1787
19.7595 | -35.5158 | -31.1859
19.7656 | -35.5163 | -31.1929
19.7717 | -35.5169 | -31.1997
19.7778 | -35.5175 | -31.2068
19.7839 | -35.518 | -31.2139
19.79 | -35.5186 | -31.2209
19.7961 | -35.5191 | -31.2279
19.8022 | -35.5197 | -31.2347
19.8083 | -35.5203 | -31.2418
19.8144 | -35.5208 | -31.2489
19.8205 | -35.5214 | -31.2561
19.8266 | -35.522 | -31.2631
19.8327 | -35.5226 | -31.2702
19.8388 | -35.5231 | -31.277
19.845 | -35.5236 | -31.284
19.8511 | -35.5242 | -31.2913
19.8572 | -35.5248 | -31.2984
19.8633 | -35.5253 | -31.3053
19.8694 | -35.5258 | -31.3126
19.8755 | -35.5264 | -31.3195
19.8816 | -35.527 | -31.3266
19.8877 | -35.5275 | -31.3337
19.8938 | -35.5281 | -31.3407
19.8999 | -35.5286 | -31.3477
19.906 | -35.5292 | -31.3548
19.9121 | -35.5297 | -31.3618
19.9182 | -35.5303 | -31.369
19.9243 | -35.5308 | -31.3761
19.9304 | -35.5314 | -31.3833
19.9365 | -35.5319 | -31.3906
19.9426 | -35.5325 | -31.3975
19.9487 | -35.5331 | -31.4045
19.9548 | -35.5336 | -31.4116
19.9609 | -35.5342 | -31.4187
19.967 | -35.5347 | -31.4258
19.9731 | -35.5353 | -31.4327
19.9792 | -35.5359 | -31.4398
19.9853 | -35.5365 | -31.447
19.9915 | -35.537 | -31.4539
19.9976 | -35.5375 | -31.4608
20.0037 | -35.5381 | -31.468
20.0098 | -35.5387 | -31.4751
20.0159 | -35.5392 | -31.482
20.022 | -35.5398 | -31.4891
20.0281 | -35.5403 | -31.4963
20.0342 | -35.5408 | -31.5034
20.0403 | -35.5414 | -31.5106
20.0464 | -35.542 | -31.5178
20.0525 | -35.5426 | -31.5244
20.0586 | -35.5432 | -31.5314
20.0647 | -35.5437 | -31.5385
20.0708 | -35.5443 | -31.5455
20.0769 | -35.5449 | -31.5527
20.083 | -35.5454 | -31.5593
20.0891 | -35.546 | -31.5663
20.0952 | -35.5465 | -31.5733
20.1013 | -35.5471 | -31.5803
20.1074 | -35.5477 | -31.5873
20.1135 | -35.5482 | -31.5946
20.1196 | -35.5488 | -31.6013
20.1257 | -35.5493 | -31.6083
20.1319 | -35.5499 | -31.6153
20.138 | -35.5505 | -31.6223
20.1441 | -35.551 | -31.6294
20.1502 | -35.5516 | -31.6365
20.1563 | -35.5521 | -31.6433
20.1624 | -35.5527 | -31.6502
20.1685 | -35.5533 | -31.6572
20.1746 | -35.5538 | -31.6643
20.1807 | -35.5544 | -31.6712
20.1868 | -35.555 | -31.6784
20.1929 | -35.5555 | -31.6849
20.199 | -35.5561 | -31.6919
20.2051 | -35.5566 | -31.699
20.2112 | -35.5571 | -31.706
20.2173 | -35.5577 | -31.7129
20.2234 | -35.5583 | -31.7197
20.2295 | -35.5588 | -31.7265
20.2356 | -35.5594 | -31.7334
20.2417 | -35.5599 | -31.7405
20.2478 | -35.5605 | -31.7474
20.2539 | -35.5611 | -31.7544
20.26 | -35.5616 | -31.7611
20.2661 | -35.5621 | -31.768
20.2723 | -35.5627 | -31.7751
20.2784 | -35.5632 | -31.7821
20.2845 | -35.5638 | -31.789
20.2906 | -35.5644 | -31.7959
20.2967 | -35.5649 | -31.8026
20.3028 | -35.5655 | -31.8095
20.3089 | -35.566 | -31.8166
20.315 | -35.5666 | -31.8235
20.3211 | -35.5672 | -31.8302
20.3272 | -35.5677 | -31.8372
20.3333 | -35.5683 | -31.8439
20.3394 | -35.5688 | -31.8509
20.3455 | -35.5694 | -31.8578
20.3516 | -35.5699 | -31.8646
20.3577 | -35.5705 | -31.8714
20.3638 | -35.571 | -31.8783
20.3699 | -35.5716 | -31.8848
20.376 | -35.5722 | -31.8917
20.3821 | -35.5727 | -31.8988
20.3882 | -35.5733 | -31.9055
20.3943 | -35.5739 | -31.9125
20.4004 | -35.5744 | -31.9191
20.4065 | -35.5749 | -31.926
20.4126 | -35.5755 | -31.9329
20.4188 | -35.5761 | -31.94
20.4249 | -35.5766 | -31.9464
20.431 | -35.5772 | -31.953
20.4371 | -35.5778 | -31.9596
20.4432 | -35.5784 | -31.9665
20.4493 | -35.5789 | -31.9734
20.4554 | -35.5795 | -31.9802
20.4615 | -35.58 | -31.9872
20.4676 | -35.5806 | -31.9944
20.4737 | -35.5812 | -32.0011
20.4798 | -35.5818 | -32.0081
20.4859 | -35.5824 | -32.0149
20.492 | -35.583 | -32.0218
20.4981 | -35.5835 | -32.0289
20.5042 | -35.5841 | -32.0359
20.5103 | -35.5846 | -32.0425
20.5164 | -35.5852 | -32.0492
20.5225 | -35.5857 | -32.0559
20.5286 | -35.5862 | -32.0627
20.5347 | -35.5868 | -32.0696
20.5408 | -35.5873 | -32.0762
20.5469 | -35.5879 | -32.083
20.553 | -35.5885 | -32.0897
20.5592 | -35.589 | -32.0965
20.5653 | -35.5896 | -32.1031
20.5714 | -35.5901 | -32.1098
20.5775 | -35.5907 | -32.1164
20.5836 | -35.5912 | -32.1235
20.5897 | -35.5918 | -32.1302
20.5958 | -35.5923 | -32.1373
20.6019 | -35.5929 | -32.1443
20.608 | -35.5934 | -32.1514
20.6141 | -35.594 | -32.158
20.6202 | -35.5945 | -32.1647
20.6263 | -35.5951 | -32.1714
20.6324 | -35.5956 | -32.1782
20.6385 | -35.5962 | -32.185
20.6446 | -35.5968 | -32.1918
20.6507 | -35.5973 | -32.1984
20.6568 | -35.5978 | -32.2054
20.6629 | -35.5984 | -32.2121
20.669 | -35.599 | -32.2189
20.6751 | -35.5996 | -32.2256
20.6812 | -35.6001 | -32.2322
20.6873 | -35.6006 | -32.2388
20.6934 | -35.6012 | -32.2457
20.6995 | -35.6018 | -32.2524
20.7057 | -35.6024 | -32.2594
20.7118 | -35.6029 | -32.2662
20.7179 | -35.6035 | -32.2723
20.724 | -35.604 | -32.2793
20.7301 | -35.6045 | -32.286
20.7362 | -35.6051 | -32.2928
20.7423 | -35.6057 | -32.2997
20.7484 | -35.6063 | -32.3065
20.7545 | -35.6068 | -32.3134
20.7606 | -35.6073 | -32.3201
20.7667 | -35.6079 | -32.3267
20.7728 | -35.6084 | -32.3333
20.7789 | -35.609 | -32.3401
20.785 | -35.6096 | -32.3468
20.7911 | -35.6101 | -32.3535
20.7972 | -35.6107 | -32.3604
20.8033 | -35.6113 | -32.3674
20.8094 | -35.6118 | -32.3743
20.8155 | -35.6124 | -32.3813
20.8216 | -35.6129 | -32.3879
20.8277 | -35.6135 | -32.3945
20.8338 | -35.614 | -32.4013
20.8399 | -35.6145 | -32.4082
20.8461 | -35.6151 | -32.4147
20.8522 | -35.6157 | -32.4212
20.8583 | -35.6162 | -32.4277
20.8644 | -35.6167 | -32.4345
20.8705 | -35.6173 | -32.4416
20.8766 | -35.6178 | -32.4484
20.8827 | -35.6184 | -32.455
20.8888 | -35.619 | -32.4617
20.8949 | -35.6195 | -32.468
20.901 | -35.62 | -32.4746
20.9071 | -35.6205 | -32.4815
20.9132 | -35.6211 | -32.4879
20.9193 | -35.6217 | -32.4946
20.9254 | -35.6222 | -32.5015
20.9315 | -35.6228 | -32.5084
20.9376 | -35.6233 | -32.5152
20.9437 | -35.6238 | -32.5221
20.9498 | -35.6244 | -32.5286
20.9559 | -35.625 | -32.5355
20.962 | -35.6255 | -32.542
20.9681 | -35.626 | -32.5483
20.9742 | -35.6265 | -32.5552
20.9803 | -35.627 | -32.5622
20.9864 | -35.6276 | -32.569
20.9926 | -35.6282 | -32.5759
20.9987 | -35.6288 | -32.5827
21.0048 | -35.6293 | -32.5894
21.0109 | -35.6299 | -32.5961
21.017 | -35.6305 | -32.6026
21.0231 | -35.631 | -32.6091
21.0292 | -35.6317 | -32.616
21.0353 | -35.6321 | -32.6222
21.0414 | -35.6327 | -32.6291
21.0475 | -35.6332 | -32.6357
21.0536 | -35.6338 | -32.6422
21.0597 | -35.6343 | -32.6488
21.0658 | -35.6348 | -32.6555
21.0719 | -35.6353 | -32.662
21.078 | -35.6358 | -32.6688
21.0841 | -35.6363 | -32.6752
21.0902 | -35.6368 | -32.6822
21.0963 | -35.6374 | -32.6889
21.1024 | -35.6379 | -32.6958
21.1085 | -35.6384 | -32.7026
21.1146 | -35.639 | -32.7093
21.1207 | -35.6395 | -32.7161
21.1268 | -35.6401 | -32.7229
21.133 | -35.6406 | -32.7294
21.1391 | -35.6412 | -32.7358
21.1452 | -35.6417 | -32.7424
21.1513 | -35.6422 | -32.7487
21.1574 | -35.6428 | -32.7557
21.1635 | -35.6434 | -32.7623
21.1696 | -35.644 | -32.7686
21.1757 | -35.6445 | -32.7748
21.1818 | -35.645 | -32.7817
21.1879 | -35.6456 | -32.7881
21.194 | -35.6461 | -32.7945
21.2001 | -35.6467 | -32.8011
21.2062 | -35.6473 | -32.8079
21.2123 | -35.6478 | -32.8138
21.2184 | -35.6484 | -32.8201
21.2245 | -35.649 | -32.8266
21.2306 | -35.6495 | -32.8332
21.2367 | -35.6501 | -32.8398
21.2428 | -35.6506 | -32.8461
21.2489 | -35.6511 | -32.8524
21.255 | -35.6517 | -32.8591
21.2611 | -35.6522 | -32.8659
21.2672 | -35.6528 | -32.8724
21.2733 | -35.6533 | -32.8791
21.2795 | -35.6539 | -32.8856
21.2856 | -35.6544 | -32.8921
21.2917 | -35.655 | -32.8984
21.2978 | -35.6556 | -32.9051
21.3039 | -35.6561 | -32.9114
21.31 | -35.6567 | -32.9184
21.3161 | -35.6572 | -32.9249
21.3222 | -35.6577 | -32.9317
21.3283 | -35.6583 | -32.9384
21.3344 | -35.6588 | -32.9449
21.3405 | -35.6593 | -32.9513
21.3466 | -35.6599 | -32.958
21.3527 | -35.6605 | -32.9639
21.3588 | -35.661 | -32.9705
21.3649 | -35.6615 | -32.9768
21.371 | -35.6621 | -32.9833
21.3771 | -35.6627 | -32.9898
21.3832 | -35.6633 | -32.9966
21.3893 | -35.6638 | -33.003
21.3954 | -35.6643 | -33.0098
21.4015 | -35.6649 | -33.0166
21.4076 | -35.6654 | -33.0231
21.4137 | -35.666 | -33.0297
21.4199 | -35.6665 | -33.0362
21.426 | -35.6671 | -33.0423
21.4321 | -35.6676 | -33.0488
21.4382 | -35.6682 | -33.0553
21.4443 | -35.6687 | -33.0617
21.4504 | -35.6692 | -33.0682
21.4565 | -35.6698 | -33.0742
21.4626 | -35.6703 | -33.0806
21.4687 | -35.6708 | -33.0876
21.4748 | -35.6714 | -33.0944
21.4809 | -35.6719 | -33.1006
21.487 | -35.6724 | -33.1072
21.4931 | -35.673 | -33.1137
21.4992 | -35.6736 | -33.1202
21.5053 | -35.6741 | -33.1266
21.5114 | -35.6746 | -33.1334
21.5175 | -35.6751 | -33.14
21.5236 | -35.6757 | -33.1467
21.5297 | -35.6762 | -33.1528
21.5358 | -35.6767 | -33.1591
21.5419 | -35.6773 | -33.1656
21.548 | -35.6778 | -33.1722
21.5541 | -35.6784 | -33.1785
21.5602 | -35.6789 | -33.1853
21.5664 | -35.6795 | -33.1914
21.5725 | -35.68 | -33.1978
21.5786 | -35.6805 | -33.2042
21.5847 | -35.6811 | -33.2108
21.5908 | -35.6816 | -33.2173
21.5969 | -35.6822 | -33.2234
21.603 | -35.6827 | -33.2301
21.6091 | -35.6832 | -33.2367
21.6152 | -35.6837 | -33.2433
21.6213 | -35.6843 | -33.2504
21.6274 | -35.6848 | -33.2571
21.6335 | -35.6854 | -33.2635
21.6396 | -35.6859 | -33.2702
21.6457 | -35.6865 | -33.2766
21.6518 | -35.687 | -33.283
21.6579 | -35.6876 | -33.2895
21.664 | -35.6882 | -33.2963
21.6701 | -35.6887 | -33.3027
21.6762 | -35.6893 | -33.3094
21.6823 | -35.6899 | -33.316
21.6884 | -35.6904 | -33.3223
21.6945 | -35.6909 | -33.3287
21.7006 | -35.6914 | -33.3352
21.7068 | -35.692 | -33.3416
21.7129 | -35.6926 | -33.3481
21.719 | -35.6931 | -33.3543
21.7251 | -35.6936 | -33.361
21.7312 | -35.6941 | -33.3673
21.7373 | -35.6947 | -33.3733
21.7434 | -35.6952 | -33.3799
21.7495 | -35.6957 | -33.3864
21.7556 | -35.6963 | -33.3929
21.7617 | -35.6969 | -33.3995
21.7678 | -35.6974 | -33.406
21.7739 | -35.698 | -33.4124
21.78 | -35.6985 | -33.4185
21.7861 | -35.6991 | -33.4252
21.7922 | -35.6996 | -33.4319
21.7983 | -35.7002 | -33.4385
21.8044 | -35.7006 | -33.4449
21.8105 | -35.7012 | -33.4511
21.8166 | -35.7018 | -33.4577
21.8227 | -35.7023 | -33.4638
21.8288 | -35.7028 | -33.47
21.8349 | -35.7034 | -33.4767
21.841 | -35.7039 | -33.4827
21.8471 | -35.7044 | -33.4888
21.8533 | -35.705 | -33.4953
21.8594 | -35.7056 | -33.5017
21.8655 | -35.7061 | -33.5082
21.8716 | -35.7066 | -33.5145
21.8777 | -35.7072 | -33.5208
21.8838 | -35.7077 | -33.5275
21.8899 | -35.7082 | -33.534
21.896 | -35.7088 | -33.5402
21.9021 | -35.7094 | -33.5468
21.9082 | -35.7099 | -33.553
21.9143 | -35.7104 | -33.559
21.9204 | -35.7109 | -33.5653
21.9265 | -35.7116 | -33.5712
21.9326 | -35.712 | -33.5775
21.9387 | -35.7126 | -33.5835
21.9448 | -35.7131 | -33.5901
21.9509 | -35.7136 | -33.5968
21.957 | -35.7142 | -33.6036
21.9631 | -35.7147 | -33.6098
21.9692 | -35.7152 | -33.6162
21.9753 | -35.7157 | -33.6229
21.9814 | -35.7162 | -33.6294
21.9875 | -35.7167 | -33.6356
21.9937 | -35.7172 | -33.642
21.9998 | -35.7178 | -33.6485
22.0059 | -35.7183 | -33.655
22.012 | -35.7189 | -33.6612
22.0181 | -35.7194 | -33.6672
22.0242 | -35.7199 | -33.6736
22.0303 | -35.7205 | -33.6801
22.0364 | -35.721 | -33.6865
22.0425 | -35.7216 | -33.6931
22.0486 | -35.7221 | -33.6999
22.0547 | -35.7227 | -33.7063
22.0608 | -35.7231 | -33.7129
22.0669 | -35.7237 | -33.7192
22.073 | -35.7242 | -33.7249
22.0791 | -35.7248 | -33.7309
22.0852 | -35.7253 | -33.7373
22.0913 | -35.7258 | -33.7436
22.0974 | -35.7264 | -33.7499
22.1035 | -35.7269 | -33.7561
22.1096 | -35.7275 | -33.7624
22.1157 | -35.728 | -33.7688
22.1218 | -35.7285 | -33.7752
22.1279 | -35.729 | -33.7811
22.134 | -35.7295 | -33.7876
22.1402 | -35.7301 | -33.7939
22.1463 | -35.7306 | -33.8001
22.1524 | -35.7311 | -33.8066
22.1585 | -35.7316 | -33.813
22.1646 | -35.7322 | -33.8188
22.1707 | -35.7327 | -33.8252
22.1768 | -35.7332 | -33.8317
22.1829 | -35.7338 | -33.8381
22.189 | -35.7343 | -33.8446
22.1951 | -35.7349 | -33.8505
22.2012 | -35.7354 | -33.857
22.2073 | -35.7359 | -33.8632
22.2134 | -35.7365 | -33.8693
22.2195 | -35.737 | -33.8755
22.2256 | -35.7376 | -33.8819
22.2317 | -35.7381 | -33.8881
22.2378 | -35.7386 | -33.8942
22.2439 | -35.7391 | -33.9007
22.25 | -35.7397 | -33.9065
22.2561 | -35.7402 | -33.9128
22.2622 | -35.7407 | -33.9188
22.2683 | -35.7413 | -33.9252
22.2744 | -35.7418 | -33.9315
22.2806 | -35.7423 | -33.9377
22.2867 | -35.7429 | -33.9441
22.2928 | -35.7434 | -33.9502
22.2989 | -35.744 | -33.9563
22.305 | -35.7445 | -33.9624
22.3111 | -35.7451 | -33.9688
22.3172 | -35.7456 | -33.9753
22.3233 | -35.7461 | -33.982
22.3294 | -35.7467 | -33.9882
22.3355 | -35.7472 | -33.9946
22.3416 | -35.7477 | -34.0008
22.3477 | -35.7482 | -34.0069
22.3538 | -35.7488 | -34.0133
22.3599 | -35.7493 | -34.0196
22.366 | -35.7498 | -34.026
22.3721 | -35.7503 | -34.0322
22.3782 | -35.7508 | -34.0384
22.3843 | -35.7513 | -34.0449
22.3904 | -35.7519 | -34.0513
22.3965 | -35.7524 | -34.0577
22.4026 | -35.7528 | -34.0639
22.4087 | -35.7534 | -34.0702
22.4148 | -35.7539 | -34.0764
22.4209 | -35.7545 | -34.0827
22.4271 | -35.755 | -34.0888
22.4332 | -35.7555 | -34.0951
22.4393 | -35.756 | -34.1013
22.4454 | -35.7566 | -34.1072
22.4515 | -35.7571 | -34.1137
22.4576 | -35.7577 | -34.1201
22.4637 | -35.7582 | -34.1266
22.4698 | -35.7588 | -34.1328
22.4759 | -35.7593 | -34.1386
22.482 | -35.7599 | -34.1449
22.4881 | -35.7604 | -34.1514
22.4942 | -35.761 | -34.1579
22.5003 | -35.7614 | -34.1642
22.5064 | -35.762 | -34.1705
22.5125 | -35.7624 | -34.1764
22.5186 | -35.763 | -34.1826
22.5247 | -35.7635 | -34.189
22.5308 | -35.764 | -34.1951
22.5369 | -35.7645 | -34.2014
22.543 | -35.765 | -34.2076
22.5491 | -35.7656 | -34.2134
22.5552 | -35.7661 | -34.2196
22.5613 | -35.7666 | -34.2259
22.5675 | -35.7672 | -34.2319
22.5736 | -35.7677 | -34.238
22.5797 | -35.7683 | -34.2443
22.5858 | -35.7688 | -34.2505
22.5919 | -35.7694 | -34.2567
22.598 | -35.7699 | -34.2634
22.6041 | -35.7704 | -34.2694
22.6102 | -35.771 | -34.2756
22.6163 | -35.7715 | -34.2819
22.6224 | -35.7721 | -34.2882
22.6285 | -35.7726 | -34.2945
22.6346 | -35.7731 | -34.3007
22.6407 | -35.7737 | -34.307
22.6468 | -35.7742 | -34.3131
22.6529 | -35.7747 | -34.3192
22.659 | -35.7752 | -34.3252
22.6651 | -35.7758 | -34.3316
22.6712 | -35.7763 | -34.3378
22.6773 | -35.7769 | -34.344
22.6834 | -35.7774 | -34.3505
22.6895 | -35.7779 | -34.3565
22.6956 | -35.7784 | -34.3627
22.7017 | -35.7789 | -34.3689
22.7079 | -35.7794 | -34.3752
22.714 | -35.78 | -34.3814
22.7201 | -35.7805 | -34.3879
22.7262 | -35.7811 | -34.3941
22.7323 | -35.7816 | -34.4003
22.7384 | -35.7822 | -34.4066
22.7445 | -35.7827 | -34.4131
22.7506 | -35.7832 | -34.4194
22.7567 | -35.7837 | -34.4256
22.7628 | -35.7843 | -34.4317
22.7689 | -35.7848 | -34.4379
22.775 | -35.7854 | -34.4441
22.7811 | -35.7859 | -34.4502
22.7872 | -35.7865 | -34.4565
22.7933 | -35.787 | -34.4625
22.7994 | -35.7875 | -34.4686
22.8055 | -35.7881 | -34.475
22.8116 | -35.7886 | -34.4812
22.8177 | -35.7892 | -34.4876
22.8238 | -35.7897 | -34.4937
22.8299 | -35.7903 | -34.4996
22.836 | -35.7908 | -34.5057
22.8421 | -35.7914 | -34.5119
22.8482 | -35.7919 | -34.5181
22.8544 | -35.7924 | -34.5244
22.8605 | -35.793 | -34.5306
22.8666 | -35.7935 | -34.5367
22.8727 | -35.794 | -34.5428
22.8788 | -35.7946 | -34.5487
22.8849 | -35.7951 | -34.5551
22.891 | -35.7957 | -34.5609
22.8971 | -35.7962 | -34.567
22.9032 | -35.7968 | -34.5728
22.9093 | -35.7973 | -34.5791
22.9154 | -35.7978 | -34.5849
22.9215 | -35.7983 | -34.5909
22.9276 | -35.7989 | -34.5971
22.9337 | -35.7994 | -34.6031
22.9398 | -35.7999 | -34.6094
22.9459 | -35.8005 | -34.6157
22.952 | -35.801 | -34.6217
22.9581 | -35.8016 | -34.6279
22.9642 | -35.8021 | -34.6342
22.9703 | -35.8026 | -34.6401
22.9764 | -35.8032 | -34.6462
22.9825 | -35.8037 | -34.6523
22.9886 | -35.8042 | -34.6583
22.9948 | -35.8047 | -34.6643
23.0009 | -35.8052 | -34.6705
23.007 | -35.8058 | -34.6768
23.0131 | -35.8063 | -34.6829
23.0192 | -35.8069 | -34.6891
23.0253 | -35.8074 | -34.6953
23.0314 | -35.8079 | -34.7016
23.0375 | -35.8085 | -34.7077
23.0436 | -35.809 | -34.7137
23.0497 | -35.8095 | -34.7201
23.0558 | -35.8101 | -34.7264
23.0619 | -35.8106 | -34.7326
23.068 | -35.8111 | -34.7389
23.0741 | -35.8116 | -34.7445
23.0802 | -35.8121 | -34.7505
23.0863 | -35.8127 | -34.7562
23.0924 | -35.8132 | -34.7624
23.0985 | -35.8137 | -34.7687
23.1046 | -35.8142 | -34.7751
23.1107 | -35.8148 | -34.7811
23.1168 | -35.8153 | -34.7871
23.1229 | -35.8158 | -34.7932
23.129 | -35.8164 | -34.7993
23.1351 | -35.8169 | -34.8051
23.1413 | -35.8174 | -34.8113
23.1474 | -35.8179 | -34.8175
23.1535 | -35.8184 | -34.8235
23.1596 | -35.8189 | -34.8297
23.1657 | -35.8195 | -34.8359
23.1718 | -35.82 | -34.8422
23.1779 | -35.8205 | -34.8483
23.184 | -35.8211 | -34.8544
23.1901 | -35.8216 | -34.8604
23.1962 | -35.8222 | -34.8665
23.2023 | -35.8227 | -34.8728
23.2084 | -35.8232 | -34.8788
23.2145 | -35.8238 | -34.8847
23.2206 | -35.8243 | -34.8906
23.2267 | -35.8248 | -34.8969
23.2328 | -35.8253 | -34.9031
23.2389 | -35.8258 | -34.9092
23.245 | -35.8263 | -34.9153
23.2511 | -35.8268 | -34.9214
23.2572 | -35.8273 | -34.9276
23.2633 | -35.8279 | -34.9336
23.2694 | -35.8284 | -34.9396
23.2755 | -35.8289 | -34.9458
23.2817 | -35.8294 | -34.9518
23.2878 | -35.8299 | -34.958
23.2939 | -35.8305 | -34.9645
23.3 | -35.831 | -34.9704
23.3061 | -35.8315 | -34.9763
23.3122 | -35.832 | -34.9826
23.3183 | -35.8325 | -34.9885
23.3244 | -35.833 | -34.9945
23.3305 | -35.8335 | -35.0008
23.3366 | -35.834 | -35.0072
23.3427 | -35.8346 | -35.0133
23.3488 | -35.8351 | -35.0195
23.3549 | -35.8356 | -35.0259
23.361 | -35.8362 | -35.0318
23.3671 | -35.8368 | -35.0383
23.3732 | -35.8373 | -35.0444
23.3793 | -35.8378 | -35.0504
23.3854 | -35.8383 | -35.0566
23.3915 | -35.8389 | -35.0627
23.3976 | -35.8394 | -35.0684
23.4037 | -35.8399 | -35.0745
23.4098 | -35.8404 | -35.0807
23.4159 | -35.841 | -35.0869
23.422 | -35.8415 | -35.0932
23.4282 | -35.842 | -35.0991
23.4343 | -35.8426 | -35.1052
23.4404 | -35.8431 | -35.1112
23.4465 | -35.8436 | -35.1172
23.4526 | -35.8441 | -35.1237
23.4587 | -35.8447 | -35.1298
23.4648 | -35.8452 | -35.1355
23.4709 | -35.8457 | -35.1417
23.477 | -35.8462 | -35.1477
23.4831 | -35.8467 | -35.154
23.4892 | -35.8473 | -35.1601
23.4953 | -35.8478 | -35.1664
23.5014 | -35.8483 | -35.1727
23.5075 | -35.8488 | -35.1785
23.5136 | -35.8493 | -35.1845
23.5197 | -35.8499 | -35.1904
23.5258 | -35.8504 | -35.1965
23.5319 | -35.8509 | -35.2023
23.538 | -35.8514 | -35.2082
23.5441 | -35.852 | -35.2146
23.5502 | -35.8525 | -35.2204
23.5563 | -35.853 | -35.2263
23.5624 | -35.8535 | -35.2326
23.5686 | -35.854 | -35.2386
23.5747 | -35.8545 | -35.2446
23.5808 | -35.855 | -35.2507
23.5869 | -35.8555 | -35.2564
23.593 | -35.8561 | -35.2623
23.5991 | -35.8566 | -35.2683
23.6052 | -35.8572 | -35.274
23.6113 | -35.8577 | -35.2801
23.6174 | -35.8582 | -35.2863
23.6235 | -35.8588 | -35.2924
23.6296 | -35.8593 | -35.2985
23.6357 | -35.8598 | -35.3045
23.6418 | -35.8603 | -35.3105
23.6479 | -35.8608 | -35.3166
23.654 | -35.8613 | -35.3226
23.6601 | -35.8618 | -35.3286
23.6662 | -35.8624 | -35.3348
23.6723 | -35.8629 | -35.3408
23.6784 | -35.8635 | -35.3469
23.6845 | -35.864 | -35.3529
23.6906 | -35.8645 | -35.3591
23.6967 | -35.865 | -35.3652
23.7028 | -35.8656 | -35.3715
23.7089 | -35.8661 | -35.3773
23.7151 | -35.8667 | -35.3834
23.7212 | -35.8672 | -35.3895
23.7273 | -35.8677 | -35.3955
23.7334 | -35.8682 | -35.4017
23.7395 | -35.8687 | -35.4076
23.7456 | -35.8693 | -35.4132
23.7517 | -35.8698 | -35.4193
23.7578 | -35.8703 | -35.4254
23.7639 | -35.8708 | -35.4314
23.77 | -35.8714 | -35.4374
23.7761 | -35.8719 | -35.4436
23.7822 | -35.8724 | -35.4494
23.7883 | -35.8729 | -35.4554
23.7944 | -35.8735 | -35.4615
23.8005 | -35.874 | -35.4677
23.8066 | -35.8745 | -35.4739
23.8127 | -35.8751 | -35.48
23.8188 | -35.8756 | -35.4861
23.8249 | -35.8762 | -35.4921
23.831 | -35.8767 | -35.4983
23.8371 | -35.8772 | -35.5042
23.8432 | -35.8777 | -35.5102
23.8493 | -35.8782 | -35.5161
23.8555 | -35.8787 | -35.5222
23.8616 | -35.8793 | -35.5283
23.8677 | -35.8798 | -35.5343
23.8738 | -35.8803 | -35.5403
23.8799 | -35.8809 | -35.5463
23.886 | -35.8814 | -35.5522
23.8921 | -35.8819 | -35.5581
23.8982 | -35.8825 | -35.5643
23.9043 | -35.883 | -35.5703
23.9104 | -35.8835 | -35.5765
23.9165 | -35.884 | -35.5825
23.9226 | -35.8845 | -35.5884
23.9287 | -35.8851 | -35.5946
23.9348 | -35.8856 | -35.6006
23.9409 | -35.8861 | -35.6067
23.947 | -35.8867 | -35.6127
23.9531 | -35.8872 | -35.6185
23.9592 | -35.8878 | -35.6245
23.9653 | -35.8883 | -35.6304
23.9714 | -35.8889 | -35.6362
23.9775 | -35.8894 | -35.6421
23.9836 | -35.89 | -35.6483
23.9897 | -35.8905 | -35.6542
23.9958 | -35.891 | -35.6602
24.002 | -35.8915 | -35.6663
24.0081 | -35.892 | -35.6722
24.0142 | -35.8926 | -35.6783
24.0203 | -35.8931 | -35.6843
24.0264 | -35.8936 | -35.6901
24.0325 | -35.8941 | -35.6962
24.0386 | -35.8947 | -35.7022
24.0447 | -35.8952 | -35.7079
24.0508 | -35.8957 | -35.7139
24.0569 | -35.8962 | -35.72
24.063 | -35.8968 | -35.7258
24.0691 | -35.8973 | -35.7315
24.0752 | -35.8979 | -35.7375
24.0813 | -35.8984 | -35.7434
24.0874 | -35.8989 | -35.7494
24.0935 | -35.8994 | -35.7554
24.0996 | -35.9 | -35.7614
24.1057 | -35.9005 | -35.7676
24.1118 | -35.901 | -35.7737
24.1179 | -35.9015 | -35.7796
24.124 | -35.902 | -35.7854
24.1301 | -35.9026 | -35.7914
24.1362 | -35.9031 | -35.7974
24.1424 | -35.9036 | -35.8036
24.1485 | -35.9042 | -35.8098
24.1546 | -35.9047 | -35.8159
24.1607 | -35.9053 | -35.8219
24.1668 | -35.9058 | -35.8276
24.1729 | -35.9063 | -35.8337
24.179 | -35.9068 | -35.8395
24.1851 | -35.9074 | -35.8455
24.1912 | -35.9078 | -35.8515
24.1973 | -35.9083 | -35.8577
24.2034 | -35.9088 | -35.8639
24.2095 | -35.9094 | -35.8698
24.2156 | -35.9098 | -35.8758
24.2217 | -35.9104 | -35.8819
24.2278 | -35.911 | -35.8878
24.2339 | -35.9115 | -35.8938
24.24 | -35.912 | -35.8997
24.2461 | -35.9126 | -35.9056
24.2522 | -35.9131 | -35.9117
24.2583 | -35.9136 | -35.9177
24.2644 | -35.9142 | -35.9239
24.2705 | -35.9147 | -35.9299
24.2766 | -35.9153 | -35.9357
24.2827 | -35.9158 | -35.9417
24.2889 | -35.9163 | -35.948
24.295 | -35.9168 | -35.9538
24.3011 | -35.9174 | -35.96
24.3072 | -35.918 | -35.9659
24.3133 | -35.9185 | -35.972
24.3194 | -35.919 | -35.9782
24.3255 | -35.9195 | -35.9845
24.3316 | -35.92 | -35.9904
24.3377 | -35.9205 | -35.9962
24.3438 | -35.921 | -36.0021
24.3499 | -35.9215 | -36.0079
24.356 | -35.922 | -36.0139
24.3621 | -35.9226 | -36.0199
24.3682 | -35.9231 | -36.0258
24.3743 | -35.9236 | -36.0318
24.3804 | -35.9241 | -36.0376
24.3865 | -35.9247 | -36.0435
24.3926 | -35.9252 | -36.0496
24.3987 | -35.9257 | -36.0557
24.4048 | -35.9262 | -36.0618
24.4109 | -35.9267 | -36.068
24.417 | -35.9273 | -36.0738
24.4231 | -35.9278 | -36.0796
24.4293 | -35.9284 | -36.0855
24.4354 | -35.9289 | -36.0917
24.4415 | -35.9295 | -36.0977
24.4476 | -35.93 | -36.1037
24.4537 | -35.9305 | -36.1096
24.4598 | -35.9311 | -36.1155
24.4659 | -35.9316 | -36.1215
24.472 | -35.9321 | -36.1276
24.4781 | -35.9326 | -36.1336
24.4842 | -35.9331 | -36.1395
24.4903 | -35.9336 | -36.1459
24.4964 | -35.9341 | -36.1516
24.5025 | -35.9346 | -36.1575
24.5086 | -35.9352 | -36.1635
24.5147 | -35.9357 | -36.1694
24.5208 | -35.9362 | -36.1755
24.5269 | -35.9368 | -36.1815
24.533 | -35.9373 | -36.1872
24.5391 | -35.9379 | -36.193
24.5452 | -35.9384 | -36.1989
24.5513 | -35.939 | -36.2045
24.5574 | -35.9395 | -36.2103
24.5635 | -35.94 | -36.2159
24.5696 | -35.9406 | -36.2218
24.5758 | -35.9411 | -36.2278
24.5819 | -35.9417 | -36.2336
24.588 | -35.9422 | -36.2394
24.5941 | -35.9427 | -36.2453
24.6002 | -35.9432 | -36.2512
24.6063 | -35.9437 | -36.2572
24.6124 | -35.9442 | -36.2634
24.6185 | -35.9448 | -36.2693
24.6246 | -35.9454 | -36.2751
24.6307 | -35.9459 | -36.2811
24.6368 | -35.9464 | -36.2872
24.6429 | -35.947 | -36.2932
24.649 | -35.9475 | -36.2992
24.6551 | -35.948 | -36.3051
24.6612 | -35.9486 | -36.311
24.6673 | -35.9491 | -36.3171
24.6734 | -35.9496 | -36.3229
24.6795 | -35.9502 | -36.3288
24.6856 | -35.9507 | -36.3346
24.6917 | -35.9512 | -36.3406
24.6978 | -35.9518 | -36.3462
24.7039 | -35.9523 | -36.3522
24.71 | -35.9527 | -36.3583
24.7162 | -35.9533 | -36.3642
24.7223 | -35.9538 | -36.3701
24.7284 | -35.9544 | -36.3758
24.7345 | -35.9549 | -36.3818
24.7406 | -35.9554 | -36.3878
24.7467 | -35.956 | -36.3939
24.7528 | -35.9565 | -36.3996
24.7589 | -35.9571 | -36.4055
24.765 | -35.9576 | -36.4112
24.7711 | -35.9581 | -36.4169
24.7772 | -35.9586 | -36.4226
24.7833 | -35.9592 | -36.4286
24.7894 | -35.9597 | -36.4344
24.7955 | -35.9602 | -36.4404
24.8016 | -35.9608 | -36.4463
24.8077 | -35.9613 | -36.4523
24.8138 | -35.9618 | -36.4582
24.8199 | -35.9623 | -36.4643
24.826 | -35.9629 | -36.4702
24.8321 | -35.9634 | -36.4762
24.8382 | -35.9639 | -36.4817
24.8443 | -35.9644 | -36.4877
24.8504 | -35.9649 | -36.4937
24.8565 | -35.9655 | -36.4997
24.8627 | -35.9661 | -36.5056
24.8688 | -35.9665 | -36.5112
24.8749 | -35.9671 | -36.5171
24.881 | -35.9676 | -36.5229
24.8871 | -35.9681 | -36.5289
24.8932 | -35.9686 | -36.535
24.8993 | -35.9691 | -36.541
24.9054 | -35.9696 | -36.5468
24.9115 | -35.9701 | -36.5526
24.9176 | -35.9706 | -36.5586
24.9237 | -35.9712 | -36.5644
24.9298 | -35.9717 | -36.5704
24.9359 | -35.9722 | -36.5766
24.942 | -35.9727 | -36.5824
24.9481 | -35.9732 | -36.5884
24.9542 | -35.9737 | -36.5945
24.9603 | -35.9742 | -36.6004
24.9664 | -35.9747 | -36.6063
24.9725 | -35.9753 | -36.6123
24.9786 | -35.9758 | -36.618
24.9847 | -35.9763 | -36.6238
24.9908 | -35.9768 | -36.6297
24.9969 | -35.9773 | -36.6358
25.0031 | -35.9778 | -36.6419
25.0092 | -35.9784 | -36.6479
25.0153 | -35.9789 | -36.6536
25.0214 | -35.9794 | -36.6597
25.0275 | -35.9799 | -36.6654
25.0336 | -35.9805 | -36.6713
25.0397 | -35.981 | -36.6773
25.0458 | -35.9816 | -36.683
25.0519 | -35.9821 | -36.6888
25.058 | -35.9826 | -36.6945
25.0641 | -35.9831 | -36.7002
25.0702 | -35.9836 | -36.706
25.0763 | -35.9841 | -36.7119
25.0824 | -35.9847 | -36.7179
25.0885 | -35.9852 | -36.7239
25.0946 | -35.9857 | -36.7298
25.1007 | -35.9862 | -36.7357
25.1068 | -35.9868 | -36.7415
25.1129 | -35.9873 | -36.7474
25.119 | -35.9878 | -36.7534
25.1251 | -35.9883 | -36.7592
25.1312 | -35.9889 | -36.7652
25.1373 | -35.9894 | -36.7712
25.1435 | -35.9899 | -36.7772
25.1496 | -35.9904 | -36.783
25.1557 | -35.9909 | -36.7887
25.1618 | -35.9914 | -36.7945
25.1679 | -35.9919 | -36.8005
25.174 | -35.9925 | -36.8065
25.1801 | -35.993 | -36.8124
25.1862 | -35.9935 | -36.8178
25.1923 | -35.994 | -36.8238
25.1984 | -35.9945 | -36.8298
25.2045 | -35.9951 | -36.8354
25.2106 | -35.9955 | -36.8415
25.2167 | -35.996 | -36.8473
25.2228 | -35.9965 | -36.853
25.2289 | -35.997 | -36.8587
25.235 | -35.9976 | -36.8642
25.2411 | -35.9981 | -36.8699
25.2472 | -35.9986 | -36.8759
25.2533 | -35.9991 | -36.8819
25.2594 | -35.9996 | -36.8875
25.2655 | -36.0001 | -36.8932
25.2716 | -36.0006 | -36.8989
25.2777 | -36.0011 | -36.9047
25.2838 | -36.0017 | -36.9107
25.29 | -36.0022 | -36.9167
25.2961 | -36.0028 | -36.9223
25.3022 | -36.0033 | -36.9281
25.3083 | -36.0038 | -36.934
25.3144 | -36.0043 | -36.9399
25.3205 | -36.0048 | -36.9459
25.3266 | -36.0053 | -36.9519
25.3327 | -36.0058 | -36.9577
25.3388 | -36.0063 | -36.9635
25.3449 | -36.0068 | -36.9693
25.351 | -36.0073 | -36.9751
25.3571 | -36.0078 | -36.9811
25.3632 | -36.0083 | -36.9867
25.3693 | -36.0088 | -36.9926
25.3754 | -36.0093 | -36.9983
25.3815 | -36.0099 | -37.0042
25.3876 | -36.0104 | -37.0099
25.3937 | -36.0109 | -37.0158
25.3998 | -36.0114 | -37.0213
25.4059 | -36.0119 | -37.0272
25.412 | -36.0125 | -37.0331
25.4181 | -36.013 | -37.0388
25.4242 | -36.0135 | -37.0447
25.4304 | -36.014 | -37.0506
25.4365 | -36.0146 | -37.0564
25.4426 | -36.0151 | -37.0624
25.4487 | -36.0156 | -37.0683
25.4548 | -36.0161 | -37.0742
25.4609 | -36.0166 | -37.08
25.467 | -36.0171 | -37.0856
25.4731 | -36.0177 | -37.0916
25.4792 | -36.0182 | -37.0976
25.4853 | -36.0187 | -37.1034
25.4914 | -36.0192 | -37.1092
25.4975 | -36.0197 | -37.1149
25.5036 | -36.0202 | -37.1207
25.5097 | -36.0208 | -37.1266
25.5158 | -36.0213 | -37.1325
25.5219 | -36.0218 | -37.1383
25.528 | -36.0223 | -37.1439
25.5341 | -36.0228 | -37.1496
25.5402 | -36.0234 | -37.1555
25.5463 | -36.0239 | -37.1613
25.5524 | -36.0244 | -37.167
25.5585 | -36.025 | -37.1727
25.5646 | -36.0255 | -37.1786
25.5707 | -36.026 | -37.1843
25.5769 | -36.0265 | -37.1901
25.583 | -36.027 | -37.196
25.5891 | -36.0275 | -37.202
25.5952 | -36.028 | -37.2079
25.6013 | -36.0285 | -37.2138
25.6074 | -36.0291 | -37.2194
25.6135 | -36.0296 | -37.2252
25.6196 | -36.0301 | -37.2312
25.6257 | -36.0307 | -37.2369
25.6318 | -36.0312 | -37.2424
25.6379 | -36.0316 | -37.2483
25.644 | -36.0321 | -37.2541
25.6501 | -36.0326 | -37.2599
25.6562 | -36.0332 | -37.2657
25.6623 | -36.0337 | -37.2716
25.6684 | -36.0342 | -37.2773
25.6745 | -36.0347 | -37.2832
25.6806 | -36.0352 | -37.289
25.6867 | -36.0357 | -37.295
25.6928 | -36.0362 | -37.3009
25.6989 | -36.0367 | -37.307
25.705 | -36.0372 | -37.3129
25.7111 | -36.0377 | -37.3188
25.7173 | -36.0382 | -37.3249
25.7234 | -36.0387 | -37.3305
25.7295 | -36.0392 | -37.3365
25.7356 | -36.0397 | -37.3422
25.7417 | -36.0402 | -37.3482
25.7478 | -36.0407 | -37.3537
25.7539 | -36.0412 | -37.3593
25.76 | -36.0417 | -37.365
25.7661 | -36.0423 | -37.3706
25.7722 | -36.0428 | -37.3766
25.7783 | -36.0433 | -37.3824
25.7844 | -36.0438 | -37.3878
25.7905 | -36.0443 | -37.3936
25.7966 | -36.0449 | -37.3994
25.8027 | -36.0454 | -37.4052
25.8088 | -36.0459 | -37.4111
25.8149 | -36.0464 | -37.4168
25.821 | -36.0469 | -37.4224
25.8271 | -36.0475 | -37.4282
25.8332 | -36.048 | -37.434
25.8393 | -36.0485 | -37.4398
25.8454 | -36.049 | -37.4455
25.8515 | -36.0495 | -37.4513
25.8576 | -36.05 | -37.4569
25.8638 | -36.0505 | -37.4626
25.8699 | -36.051 | -37.4681
25.876 | -36.0515 | -37.4738
25.8821 | -36.052 | -37.4794
25.8882 | -36.0525 | -37.485
25.8943 | -36.053 | -37.4909
25.9004 | -36.0535 | -37.4969
25.9065 | -36.0541 | -37.5028
25.9126 | -36.0545 | -37.5083
25.9187 | -36.055 | -37.514
25.9248 | -36.0555 | -37.5197
25.9309 | -36.0561 | -37.5256
25.937 | -36.0565 | -37.5314
25.9431 | -36.057 | -37.5371
25.9492 | -36.0575 | -37.5426
25.9553 | -36.058 | -37.5485
25.9614 | -36.0585 | -37.554
25.9675 | -36.059 | -37.5596
25.9736 | -36.0596 | -37.5654
25.9797 | -36.0601 | -37.5712
25.9858 | -36.0605 | -37.5771
25.9919 | -36.061 | -37.5831
25.998 | -36.0615 | -37.5888
26.0042 | -36.062 | -37.5943
26.0103 | -36.0625 | -37.6001
26.0164 | -36.063 | -37.606
26.0225 | -36.0635 | -37.6116
26.0286 | -36.0641 | -37.6173
26.0347 | -36.0645 | -37.6227
26.0408 | -36.065 | -37.6285
26.0469 | -36.0655 | -37.634
26.053 | -36.066 | -37.6397
26.0591 | -36.0665 | -37.6454
26.0652 | -36.0669 | -37.651
26.0713 | -36.0674 | -37.6567
26.0774 | -36.0679 | -37.6624
26.0835 | -36.0684 | -37.6682
26.0896 | -36.0689 | -37.6741
26.0957 | -36.0694 | -37.68
26.1018 | -36.0699 | -37.6854
26.1079 | -36.0704 | -37.6912
26.114 | -36.0709 | -37.6969
26.1201 | -36.0713 | -37.7027
26.1262 | -36.0719 | -37.7084
26.1323 | -36.0724 | -37.7142
26.1384 | -36.0729 | -37.7199
26.1445 | -36.0734 | -37.7257
26.1507 | -36.0739 | -37.7313
26.1568 | -36.0744 | -37.7369
26.1629 | -36.0749 | -37.7428
26.169 | -36.0753 | -37.7485
26.1751 | -36.0758 | -37.754
26.1812 | -36.0763 | -37.7595
26.1873 | -36.0768 | -37.7653
26.1934 | -36.0773 | -37.771
26.1995 | -36.0777 | -37.7765
26.2056 | -36.0782 | -37.7823
26.2117 | -36.0787 | -37.7881
26.2178 | -36.0792 | -37.7938
26.2239 | -36.0797 | -37.7996
26.23 | -36.0802 | -37.8055
26.2361 | -36.0807 | -37.8112
26.2422 | -36.0812 | -37.8166
26.2483 | -36.0816 | -37.8223
26.2544 | -36.0821 | -37.8279
26.2605 | -36.0827 | -37.8338
26.2666 | -36.0832 | -37.8393
26.2727 | -36.0837 | -37.8451
26.2788 | -36.0842 | -37.8506
26.2849 | -36.0847 | -37.8565
26.2911 | -36.0852 | -37.8623
26.2972 | -36.0857 | -37.8679
26.3033 | -36.0862 | -37.8737
26.3094 | -36.0867 | -37.8793
26.3155 | -36.0872 | -37.8849
26.3216 | -36.0876 | -37.8907
26.3277 | -36.0881 | -37.8963
26.3338 | -36.0887 | -37.9016
26.3399 | -36.0891 | -37.9074
26.346 | -36.0896 | -37.9126
26.3521 | -36.0901 | -37.9186
26.3582 | -36.0906 | -37.9242
26.3643 | -36.0911 | -37.9297
26.3704 | -36.0916 | -37.9355
26.3765 | -36.0921 | -37.9411
26.3826 | -36.0926 | -37.9465
26.3887 | -36.0931 | -37.9522
26.3948 | -36.0936 | -37.9581
26.4009 | -36.0941 | -37.9637
26.407 | -36.0946 | -37.9693
26.4131 | -36.0951 | -37.9751
26.4192 | -36.0956 | -37.9809
26.4253 | -36.0961 | -37.9867
26.4314 | -36.0966 | -37.9925
26.4376 | -36.0971 | -37.998
26.4437 | -36.0976 | -38.0038
26.4498 | -36.0981 | -38.0095
26.4559 | -36.0985 | -38.0147
26.462 | -36.099 | -38.0204
26.4681 | -36.0995 | -38.0258
26.4742 | -36.1 | -38.0314
26.4803 | -36.1005 | -38.0369
26.4864 | -36.101 | -38.0424
26.4925 | -36.1015 | -38.0481
26.4986 | -36.1019 | -38.054
26.5047 | -36.1024 | -38.0596
26.5108 | -36.1029 | -38.0653
26.5169 | -36.1034 | -38.071
26.523 | -36.1038 | -38.0765
26.5291 | -36.1044 | -38.0822
26.5352 | -36.1048 | -38.0877
26.5413 | -36.1052 | -38.0934
26.5474 | -36.1058 | -38.0995
26.5535 | -36.1062 | -38.1055
26.5596 | -36.1067 | -38.1109
26.5657 | -36.1071 | -38.1165
26.5718 | -36.1076 | -38.1222
26.578 | -36.1081 | -38.1279
26.5841 | -36.1086 | -38.1335
26.5902 | -36.1091 | -38.1393
26.5963 | -36.1096 | -38.1446
26.6024 | -36.1101 | -38.1503
26.6085 | -36.1106 | -38.1561
26.6146 | -36.1111 | -38.1618
26.6207 | -36.1115 | -38.1676
26.6268 | -36.112 | -38.1733
26.6329 | -36.1125 | -38.1788
26.639 | -36.113 | -38.1845
26.6451 | -36.1135 | -38.1901
26.6512 | -36.1139 | -38.1956
26.6573 | -36.1144 | -38.2013
26.6634 | -36.1148 | -38.2069
26.6695 | -36.1153 | -38.2129
26.6756 | -36.1158 | -38.2187
26.6817 | -36.1162 | -38.2244
26.6878 | -36.1167 | -38.23
26.6939 | -36.1172 | -38.2357
26.7 | -36.1176 | -38.2414
26.7061 | -36.1181 | -38.2467
26.7122 | -36.1186 | -38.2522
26.7183 | -36.1191 | -38.2578
26.7245 | -36.1195 | -38.2636
26.7306 | -36.12 | -38.2692
26.7367 | -36.1204 | -38.2748
26.7428 | -36.1209 | -38.2801
26.7489 | -36.1215 | -38.2857
26.755 | -36.1219 | -38.2909
26.7611 | -36.1224 | -38.2963
26.7672 | -36.1229 | -38.3019
26.7733 | -36.1233 | -38.3075
26.7794 | -36.1238 | -38.3129
26.7855 | -36.1243 | -38.3185
26.7916 | -36.1248 | -38.3245
26.7977 | -36.1253 | -38.3297
26.8038 | -36.1257 | -38.3355
26.8099 | -36.1262 | -38.3411
26.816 | -36.1267 | -38.3469
26.8221 | -36.1271 | -38.3524
26.8282 | -36.1276 | -38.3576
26.8343 | -36.128 | -38.3632
26.8404 | -36.1285 | -38.3684
26.8465 | -36.129 | -38.3745
26.8526 | -36.1294 | -38.3803
26.8587 | -36.1299 | -38.3856
26.8649 | -36.1304 | -38.391
26.871 | -36.1309 | -38.3963
26.8771 | -36.1313 | -38.4017
26.8832 | -36.1318 | -38.4073
26.8893 | -36.1322 | -38.4128
26.8954 | -36.1327 | -38.4179
26.9015 | -36.1332 | -38.4235
26.9076 | -36.1336 | -38.4292
26.9137 | -36.1341 | -38.4347
26.9198 | -36.1345 | -38.4402
26.9259 | -36.135 | -38.4461
26.932 | -36.1355 | -38.4515
26.9381 | -36.1359 | -38.4571
26.9442 | -36.1363 | -38.4628
26.9503 | -36.1368 | -38.4683
26.9564 | -36.1373 | -38.4737
26.9625 | -36.1378 | -38.4793
26.9686 | -36.1382 | -38.485
26.9747 | -36.1387 | -38.4907
26.9808 | -36.1391 | -38.4959
26.9869 | -36.1396 | -38.5014
26.993 | -36.14 | -38.5071
26.9991 | -36.1405 | -38.5127
27.0052 | -36.141 | -38.5185
27.0114 | -36.1415 | -38.5241
27.0175 | -36.142 | -38.5292
27.0236 | -36.1425 | -38.535
27.0297 | -36.143 | -38.5405
27.0358 | -36.1435 | -38.5463
27.0419 | -36.144 | -38.552
27.048 | -36.1445 | -38.5575
27.0541 | -36.1449 | -38.5627
27.0602 | -36.1454 | -38.5682
27.0663 | -36.1458 | -38.574
27.0724 | -36.1463 | -38.5795
27.0785 | -36.1468 | -38.585
27.0846 | -36.1472 | -38.5902
27.0907 | -36.1476 | -38.5958
27.0968 | -36.1481 | -38.6016
27.1029 | -36.1485 | -38.6071
27.109 | -36.149 | -38.6125
27.1151 | -36.1495 | -38.6179
27.1212 | -36.1499 | -38.6232
27.1273 | -36.1504 | -38.6288
27.1334 | -36.1509 | -38.6343
27.1395 | -36.1513 | -38.6398
27.1456 | -36.1518 | -38.6457
27.1518 | -36.1522 | -38.6514
27.1579 | -36.1526 | -38.6566
27.164 | -36.153 | -38.6622
27.1701 | -36.1535 | -38.6676
27.1762 | -36.1539 | -38.6729
27.1823 | -36.1544 | -38.6783
27.1884 | -36.1548 | -38.6839
27.1945 | -36.1553 | -38.6888
27.2006 | -36.1557 | -38.6946
27.2067 | -36.1562 | -38.7
27.2128 | -36.1567 | -38.7054
27.2189 | -36.1571 | -38.7111
27.225 | -36.1575 | -38.7165
27.2311 | -36.158 | -38.7212
27.2372 | -36.1584 | -38.7267
27.2433 | -36.1589 | -38.7318
27.2494 | -36.1594 | -38.7373
27.2555 | -36.1599 | -38.7429
27.2616 | -36.1604 | -38.748
27.2677 | -36.1608 | -38.7537
27.2738 | -36.1612 | -38.7593
27.2799 | -36.1617 | -38.7645
27.286 | -36.1621 | -38.7699
27.2921 | -36.1626 | -38.7754
27.2983 | -36.1631 | -38.7806
27.3044 | -36.1635 | -38.7861
27.3105 | -36.164 | -38.7919
27.3166 | -36.1645 | -38.7973
27.3227 | -36.1649 | -38.8026
27.3288 | -36.1654 | -38.8082
27.3349 | -36.1658 | -38.8138
27.341 | -36.1663 | -38.8192
27.3471 | -36.1667 | -38.8245
27.3532 | -36.1672 | -38.83
27.3593 | -36.1676 | -38.8353
27.3654 | -36.1681 | -38.8406
27.3715 | -36.1686 | -38.8459
27.3776 | -36.169 | -38.8514
27.3837 | -36.1695 | -38.8566
27.3898 | -36.1699 | -38.862
27.3959 | -36.1703 | -38.8676
27.402 | -36.1708 | -38.8727
27.4081 | -36.1712 | -38.8783
27.4142 | -36.1717 | -38.8838
27.4203 | -36.1722 | -38.8893
27.4264 | -36.1727 | -38.8945
27.4325 | -36.1731 | -38.8996
27.4387 | -36.1736 | -38.9055
27.4448 | -36.174 | -38.9106
27.4509 | -36.1744 | -38.9163
27.457 | -36.1749 | -38.9221
27.4631 | -36.1753 | -38.9274
27.4692 | -36.1758 | -38.9331
27.4753 | -36.1761 | -38.9382
27.4814 | -36.1766 | -38.9438
27.4875 | -36.1771 | -38.9497
27.4936 | -36.1775 | -38.9552
27.4997 | -36.178 | -38.9606
27.5058 | -36.1785 | -38.9662
27.5119 | -36.1789 | -38.9715
27.518 | -36.1793 | -38.9772
27.5241 | -36.1797 | -38.9829
27.5302 | -36.1802 | -38.9889
27.5363 | -36.1806 | -38.9944
27.5424 | -36.181 | -38.9996
27.5485 | -36.1815 | -39.005
27.5546 | -36.1819 | -39.0105
27.5607 | -36.1824 | -39.0158
27.5668 | -36.1828 | -39.0211
27.5729 | -36.1833 | -39.0266
27.5791 | -36.1837 | -39.0317
27.5852 | -36.1842 | -39.0372
27.5913 | -36.1847 | -39.0426
27.5974 | -36.185 | -39.0482
27.6035 | -36.1855 | -39.0539
27.6096 | -36.1859 | -39.0598
27.6157 | -36.1863 | -39.0646
27.6218 | -36.1868 | -39.0697
27.6279 | -36.1872 | -39.0752
27.634 | -36.1877 | -39.0804
27.6401 | -36.1881 | -39.0858
27.6462 | -36.1886 | -39.0911
27.6523 | -36.1891 | -39.0964
27.6584 | -36.1895 | -39.102
27.6645 | -36.1899 | -39.1075
27.6706 | -36.1904 | -39.1132
27.6767 | -36.1908 | -39.1183
27.6828 | -36.1912 | -39.1231
27.6889 | -36.1917 | -39.1286
27.695 | -36.1921 | -39.134
27.7011 | -36.1926 | -39.1393
27.7072 | -36.193 | -39.1445
27.7133 | -36.1934 | -39.15
27.7194 | -36.1939 | -39.1549
27.7256 | -36.1943 | -39.1604
27.7317 | -36.1947 | -39.1659
27.7378 | -36.1951 | -39.1712
27.7439 | -36.1956 | -39.177
27.75 | -36.1961 | -39.1824
27.7561 | -36.1965 | -39.1877
27.7622 | -36.197 | -39.1934
27.7683 | -36.1974 | -39.1988
27.7744 | -36.1979 | -39.2043
27.7805 | -36.1983 | -39.2098
27.7866 | -36.1987 | -39.2151
27.7927 | -36.1991 | -39.2203
27.7988 | -36.1996 | -39.2256
27.8049 | -36.2 | -39.2312
27.811 | -36.2004 | -39.2367
27.8171 | -36.2009 | -39.2421
27.8232 | -36.2013 | -39.2473
27.8293 | -36.2018 | -39.2525
27.8354 | -36.2022 | -39.2579
27.8415 | -36.2027 | -39.2634
27.8476 | -36.2031 | -39.2681
27.8537 | -36.2035 | -39.2731
27.8598 | -36.2039 | -39.2781
27.866 | -36.2044 | -39.2834
27.8721 | -36.2049 | -39.2888
27.8782 | -36.2053 | -39.2945
27.8843 | -36.2057 | -39.3002
27.8904 | -36.2061 | -39.3055
27.8965 | -36.2065 | -39.3109
27.9026 | -36.2069 | -39.3163
27.9087 | -36.2073 | -39.3217
27.9148 | -36.2077 | -39.3269
27.9209 | -36.2081 | -39.3321
27.927 | -36.2085 | -39.3378
27.9331 | -36.2089 | -39.3427
27.9392 | -36.2094 | -39.348
27.9453 | -36.2098 | -39.3533
27.9514 | -36.2101 | -39.3584
27.9575 | -36.2105 | -39.364
27.9636 | -36.2109 | -39.3692
27.9697 | -36.2114 | -39.3743
27.9758 | -36.2119 | -39.3799
27.9819 | -36.2123 | -39.3853
27.988 | -36.2128 | -39.3906
27.9941 | -36.2132 | -39.3959
28.0002 | -36.2135 | -39.4014
28.0063 | -36.2139 | -39.4069
28.0125 | -36.2144 | -39.4122
28.0186 | -36.2149 | -39.4179
28.0247 | -36.2153 | -39.4227
28.0308 | -36.2158 | -39.428
28.0369 | -36.2163 | -39.4334
28.043 | -36.2167 | -39.4389
28.0491 | -36.2172 | -39.4443
28.0552 | -36.2176 | -39.4495
28.0613 | -36.218 | -39.4548
28.0674 | -36.2185 | -39.4605
28.0735 | -36.2189 | -39.4663

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.9272 | -44.1205
21.4908 | -35.858 | -36.5474
16.4948 | -35.1724 | -28.9529
12.6602 | -34.7695 | -22.4923
9.7171 | -34.5202 | -17.4278
7.4582 | -34.351 | -13.334
5.7244 | -34.2305 | -10.0032
4.3936 | -34.1523 | -7.3272
3.3722 | -34.1204 | -5.2088
2.5883 | -34.1023 | -3.5982
1.9866 | -34.0896 | -2.3136
1.5248 | -34.0865 | -1.2771
1.1703 | -34.0942 | -0.59459
1 | -34.0865 | -0.26014
0.89822 | -34.0762 | 0.02405
0.68941 | -34.0656 | 0.62909
0.52919 | -34.0527 | 1.096
0.40612 | -34.0427 | 1.7227
0.31171 | -34.046 | 2.1742
0.23928 | -34.0445 | 2.9098
0.18358 | -34.0465 | 3.3379
0.14091 | -34.0579 | 4.2191
0.10819 | -34.0543 | 5.3912
0.083017 | -34.0867 | 6.8728
0.063705 | -34.128 | 8.7477
0.048923 | -34.1848 | 11.2402
0.037479 | -34.2939 | 14.4423
0.028801 | -34.467 | 18.4829
0.022078 | -34.7795 | 23.889
0.017023 | -35.2589 | 30.3178
0.013018 | -36.0509 | 37.9756
0.010014 | -37.1639 | 45.8985

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.00378 | -1.9627
0.007559 | -29.755
0.011339 | -36.1915
0.015119 | -40.2483
0.018898 | -43.2731
0.022678 | -45.9633
0.026458 | -46.7499
0.030237 | -47.8085
0.034017 | -49.0809
0.037796 | -49.9439
0.041576 | -51.3819
0.045356 | -52.7109
0.049135 | -53.3688
0.052915 | -54.7402
0.056695 | -55.7221
0.060474 | -56.7177
0.064254 | -57.5117
0.068034 | -57.422
0.071813 | -58.5326
0.075593 | -59.5292
0.079373 | -59.916
0.083152 | -59.9211
0.086932 | -59.7099
0.090712 | -60.0704
0.094491 | -61.3566
0.098271 | -61.6144
0.10205 | -61.3521
0.10583 | -62.3838
0.10961 | -62.5735
0.11339 | -64.3361
0.11717 | -64.4541
0.12095 | -64.591
0.12473 | -65.1773
0.12851 | -66.3089
0.13229 | -65.7696
0.13607 | -66.1779
0.13985 | -65.8922
0.14363 | -66.6094
0.14741 | -65.6343
0.15119 | -66.5125
0.15497 | -66.5331
0.15874 | -66.3144
0.16253 | -66.6734
0.1663 | -67.0533
0.17008 | -68.0824
0.17386 | -69.4659
0.17764 | -69.3424
0.18142 | -69.4073
0.1852 | -68.6437
0.18898 | -69.0092
0.19276 | -69.3262
0.19654 | -70.4975
0.20032 | -72.0339
0.2041 | -71.798
0.20788 | -70.96
0.21166 | -70.7143
0.21544 | -71.0604
0.21922 | -71.3082
0.223 | -71.215
0.22678 | -70.8478
0.23056 | -70.5321
0.23434 | -71.4867
0.23812 | -72.2525
0.2419 | -71.636
0.24568 | -72.3045
0.24946 | -73.114
0.25324 | -73.12
0.25702 | -72.3168
0.2608 | -72.5978
0.26458 | -72.7348
0.26836 | -73.3479
0.27214 | -72.6988
0.27591 | -72.6878
0.27969 | -74.1318
0.28347 | -73.9693
0.28725 | -73.4041
0.29103 | -73.9206
0.29481 | -75.0306
0.29859 | -75.6562
0.30237 | -74.2437
0.30615 | -74.074
0.30993 | -74.4329
0.31371 | -74.9525
0.31749 | -75.1863
0.32127 | -75.4534
0.32505 | -76.4859
0.32883 | -75.6783
0.33261 | -75.2894
0.33639 | -74.7013
0.34017 | -74.9356
0.34395 | -75.4672
0.34773 | -75.2366
0.35151 | -76.3725
0.35529 | -77.0743
0.35907 | -76.3935
0.36285 | -76.8637
0.36663 | -77.5939
0.3704 | -76.8451
0.37418 | -76.3683
0.37796 | -76.7581
0.38174 | -76.7321
0.38552 | -76.0291
0.3893 | -76.2654
0.39308 | -77.3995
0.39686 | -77.5043
0.40064 | -77.5749
0.40442 | -77.3701
0.4082 | -77.9506
0.41198 | -78.5602
0.41576 | -79.1315
0.41954 | -78.6909
0.42332 | -78.0922
0.4271 | -77.5526
0.43088 | -78.4912
0.43466 | -78.4705
0.43844 | -78.4586
0.44222 | -78.7818
0.446 | -78.3936
0.44978 | -77.7784
0.45356 | -77.8404
0.45734 | -78.2875
0.46112 | -78.3936
0.4649 | -77.5593
0.46868 | -78.594
0.47246 | -79.7826
0.47624 | -80.072
0.48002 | -79.0553
0.48379 | -78.5338
0.48757 | -79.1287
0.49135 | -78.8985
0.49513 | -78.7734
0.49891 | -79.2285
0.50269 | -79.2356
0.50647 | -79.5227
0.51025 | -79.5639
0.51403 | -79.1816
0.51781 | -79.4946
0.52159 | -79.3608
0.52537 | -79.2408
0.52915 | -79.5862
0.53293 | -79.8507
0.53671 | -79.6998
0.54049 | -79.2742
0.54427 | -80.3212
0.54805 | -80.9829
0.55183 | -81.141
0.55561 | -81.3836
0.55939 | -81.7333
0.56317 | -81.6584
0.56695 | -80.9934
0.57073 | -80.7659
0.57451 | -80.779
0.57829 | -80.5401
0.58207 | -80.2553
0.58584 | -80.0265
0.58962 | -79.962
0.5934 | -80.4492
0.59718 | -81.1867
0.60096 | -81.186
0.60474 | -80.9014
0.60852 | -80.9031
0.6123 | -80.9012
0.61608 | -80.9693
0.61986 | -81.0517
0.62364 | -80.7536
0.62742 | -80.431
0.6312 | -80.6359
0.63498 | -81.6423
0.63876 | -82.6584
0.64254 | -82.3635
0.64632 | -82.006
0.6501 | -82.0977
0.65388 | -82.4445
0.65766 | -82.4471
0.66144 | -82.0001
0.66522 | -81.9369
0.669 | -82.0796
0.67278 | -82.0832
0.67656 | -82.3378
0.68034 | -82.6343
0.68412 | -82.2779
0.6879 | -82.1359
0.69168 | -82.7061
0.69546 | -82.8695
0.69924 | -82.4557
0.70301 | -82.1452
0.70679 | -81.956
0.71057 | -82.13
0.71435 | -82.6546
0.71813 | -82.882
0.72191 | -82.9582
0.72569 | -83.2684
0.72947 | -83.7774
0.73325 | -84.3344
0.73703 | -84.0778
0.74081 | -83.1455
0.74459 | -82.4821
0.74837 | -82.465
0.75215 | -82.8439
0.75593 | -82.9187
0.75971 | -83.0367
0.76349 | -83.2965
0.76727 | -83.5788
0.77105 | -83.8996
0.77483 | -84.2677
0.77861 | -84.3385
0.78239 | -84.4336
0.78617 | -84.2546
0.78995 | -84.2662
0.79373 | -84.5123
0.79751 | -84.4445
0.80129 | -84.5367
0.80507 | -84.812
0.80884 | -85.0079
0.81262 | -85.3488
0.8164 | -85.2775
0.82018 | -84.7776
0.82396 | -84.629
0.82774 | -84.7979
0.83152 | -84.7157
0.8353 | -84.3073
0.83908 | -84.1479
0.84286 | -84.6017
0.84664 | -85.199
0.85042 | -85.1585
0.8542 | -84.7042
0.85798 | -84.4173
0.86176 | -84.1982
0.86554 | -84.0144
0.86932 | -83.8248
0.8731 | -83.5555
0.87688 | -83.3791
0.88066 | -83.6228
0.88444 | -84.0157
0.88822 | -84.2967
0.892 | -84.7328
0.89578 | -85.3215
0.89956 | -85.6318
0.90334 | -85.4486
0.90712 | -85.0141
0.9109 | -85.0065
0.91467 | -85.2521
0.91845 | -85.6232
0.92223 | -85.9759
0.92601 | -86.1572
0.92979 | -85.9337
0.93357 | -85.9122
0.93735 | -86.2065
0.94113 | -86.1367
0.94491 | -85.8866
0.94869 | -85.656
0.95247 | -85.72
0.95625 | -86.238
0.96003 | -86.7455
0.96381 | -86.3076
0.96759 | -85.6659
0.97137 | -85.5911
0.97515 | -85.817
0.97893 | -85.9087
0.98271 | -85.5882
0.98649 | -85.5504
0.99027 | -86.1918
0.99405 | -86.6605
0.99783 | -86.5986
1.0016 | -86.3745
1.0054 | -86.3153
1.0092 | -86.2417
1.0129 | -86.1348
1.0167 | -86.3671
1.0205 | -86.7785
1.0243 | -87.0032
1.0281 | -87.171
1.0318 | -87.1869
1.0356 | -86.9184
1.0394 | -87.0175
1.0432 | -87.1565
1.047 | -86.9059
1.0507 | -86.558
1.0545 | -86.4837
1.0583 | -86.3487
1.0621 | -86.104
1.0659 | -86.0242
1.0696 | -85.7302
1.0734 | -85.3906
1.0772 | -85.5516
1.081 | -85.9226
1.0848 | -85.9678
1.0885 | -86.2838
1.0923 | -86.7626
1.0961 | -87.2321
1.0999 | -87.4623
1.1037 | -87.4812
1.1074 | -87.4173
1.1112 | -87.4166
1.115 | -87.4827
1.1188 | -87.5261
1.1226 | -87.611
1.1263 | -87.7608
1.1301 | -87.8572
1.1339 | -87.8738
1.1377 | -87.8648
1.1415 | -87.799
1.1452 | -87.5411
1.149 | -87.1919
1.1528 | -87.1911
1.1566 | -87.6377
1.1604 | -88.054
1.1641 | -88.1148
1.1679 | -88.0803
1.1717 | -88.0485
1.1755 | -87.8091
1.1792 | -87.4539
1.183 | -87.3052
1.1868 | -87.3351
1.1906 | -87.3617
1.1944 | -87.5319
1.1981 | -88.0243
1.2019 | -88.5512
1.2057 | -88.6557
1.2095 | -88.3902
1.2133 | -88.1849
1.217 | -88.1901
1.2208 | -88.1052
1.2246 | -87.7325
1.2284 | -87.3359
1.2322 | -87.1782
1.2359 | -87.3278
1.2397 | -87.6468
1.2435 | -87.9079
1.2473 | -87.9687
1.2511 | -87.8764
1.2548 | -87.864
1.2586 | -88.0165
1.2624 | -88.2281
1.2662 | -88.4679
1.27 | -88.73
1.2737 | -88.9736
1.2775 | -89.2309
1.2813 | -89.5273
1.2851 | -89.67
1.2889 | -89.4021
1.2926 | -88.9759
1.2964 | -88.774
1.3002 | -88.853
1.304 | -89.0789
1.3078 | -89.3103
1.3115 | -89.4076
1.3153 | -89.3337
1.3191 | -89.1341
1.3229 | -88.8985
1.3267 | -88.7004
1.3304 | -88.5464
1.3342 | -88.5078
1.338 | -88.6504
1.3418 | -88.821
1.3456 | -88.8814
1.3493 | -88.9643
1.3531 | -89.1693
1.3569 | -89.5116
1.3607 | -89.9817
1.3645 | -90.2658
1.3682 | -90.0496
1.372 | -89.604
1.3758 | -89.3478
1.3796 | -89.4575
1.3834 | -89.7251
1.3871 | -89.7519
1.3909 | -89.4947
1.3947 | -89.2074
1.3985 | -89.1377
1.4022 | -89.3871
1.406 | -89.8259
1.4098 | -90.21
1.4136 | -90.3362
1.4174 | -90.1667
1.4211 | -89.8328
1.4249 | -89.5059
1.4287 | -89.3304
1.4325 | -89.365
1.4363 | -89.584
1.44 | -89.9134
1.4438 | -90.1866
1.4476 | -90.2067
1.4514 | -89.9311
1.4552 | -89.493
1.4589 | -89.0811
1.4627 | -88.8547
1.4665 | -88.8414
1.4703 | -89.0653
1.4741 | -89.4891
1.4778 | -89.8099
1.4816 | -89.7565
1.4854 | -89.4584
1.4892 | -89.2039
1.493 | -89.1754
1.4967 | -89.3806
1.5005 | -89.7292
1.5043 | -90.0228
1.5081 | -90.0734
1.5119 | -89.9963
1.5156 | -90.1314
1.5194 | -90.5209
1.5232 | -90.7553
1.527 | -90.6628
1.5308 | -90.5914
1.5345 | -90.7062
1.5383 | -90.8649
1.5421 | -90.9373
1.5459 | -90.9656
1.5497 | -91.0296
1.5534 | -91.1782
1.5572 | -91.2819
1.561 | -91.115
1.5648 | -90.8604
1.5686 | -90.7626
1.5723 | -90.7815
1.5761 | -90.9187
1.5799 | -91.1735
1.5837 | -91.3637
1.5875 | -91.394
1.5912 | -91.3743
1.595 | -91.2979
1.5988 | -91.041
1.6026 | -90.7639
1.6063 | -90.5916
1.6101 | -90.4545
1.6139 | -90.5163
1.6177 | -90.8529
1.6215 | -91.0932
1.6252 | -90.9933
1.629 | -90.7724
1.6328 | -90.7935
1.6366 | -90.8269
1.6404 | -90.8938
1.6441 | -90.9005
1.6479 | -90.8299
1.6517 | -90.7447
1.6555 | -90.7183
1.6593 | -90.7921
1.663 | -90.9631
1.6668 | -91.1634
1.6706 | -91.3186
1.6744 | -91.3625
1.6782 | -91.2955
1.6819 | -91.1684
1.6857 | -91.0617
1.6895 | -91.0321
1.6933 | -91.0354
1.6971 | -90.9934
1.7008 | -90.8986
1.7046 | -90.8123
1.7084 | -90.793
1.7122 | -90.86
1.716 | -90.9974
1.7197 | -91.1585
1.7235 | -91.2737
1.7273 | -91.2962
1.7311 | -91.2185
1.7349 | -91.089
1.7386 | -91.0292
1.7424 | -91.1255
1.7462 | -91.3492
1.75 | -91.6142
1.7538 | -91.8545
1.7575 | -92.032
1.7613 | -92.1121
1.7651 | -92.1111
1.7689 | -92.1014
1.7727 | -92.0655
1.7764 | -91.9237
1.7802 | -91.6347
1.784 | -91.3122
1.7878 | -91.1273
1.7916 | -91.148
1.7953 | -91.3144
1.7991 | -91.5148
1.8029 | -91.6994
1.8067 | -91.8368
1.8105 | -91.8649
1.8142 | -91.7914
1.818 | -91.6578
1.8218 | -91.5093
1.8256 | -91.4035
1.8293 | -91.3984
1.8331 | -91.4494
1.8369 | -91.4441
1.8407 | -91.3887
1.8445 | -91.4076
1.8482 | -91.5708
1.852 | -91.8562
1.8558 | -92.1499
1.8596 | -92.3339
1.8634 | -92.3712
1.8671 | -92.3484
1.8709 | -92.3226
1.8747 | -92.2261
1.8785 | -92.0086
1.8823 | -91.7711
1.886 | -91.656
1.8898 | -91.7566
1.8936 | -92.0232
1.8974 | -92.2892
1.9012 | -92.4075
1.9049 | -92.3646
1.9087 | -92.271
1.9125 | -92.209
1.9163 | -92.2015
1.9201 | -92.2374
1.9238 | -92.2894
1.9276 | -92.38
1.9314 | -92.5055
1.9352 | -92.5984
1.939 | -92.5877
1.9427 | -92.4654
1.9465 | -92.3259
1.9503 | -92.2369
1.9541 | -92.2157
1.9579 | -92.288
1.9616 | -92.4398
1.9654 | -92.6097
1.9692 | -92.712
1.973 | -92.7517
1.9768 | -92.7645
1.9805 | -92.7186
1.9843 | -92.6218
1.9881 | -92.5463
1.9919 | -92.5139
1.9957 | -92.4351
1.9994 | -92.2672
2.0032 | -92.1435
2.007 | -92.2094
2.0108 | -92.4277
2.0146 | -92.613
2.0183 | -92.6063
2.0221 | -92.4481
2.0259 | -92.2669
2.0297 | -92.1583
2.0335 | -92.1524
2.0372 | -92.2355
2.041 | -92.3461
2.0448 | -92.4075
2.0486 | -92.4028
2.0523 | -92.4169
2.0561 | -92.5355
2.0599 | -92.7528
2.0637 | -92.9523
2.0675 | -92.978
2.0712 | -92.8078
2.075 | -92.5658
2.0788 | -92.4269
2.0826 | -92.4843
2.0864 | -92.734
2.0901 | -93.0797
2.0939 | -93.3997
2.0977 | -93.6204
2.1015 | -93.7405
2.1053 | -93.7617
2.109 | -93.6973
2.1128 | -93.587
2.1166 | -93.4854
2.1204 | -93.4232
2.1242 | -93.4055
2.1279 | -93.4035
2.1317 | -93.3937
2.1355 | -93.3643
2.1393 | -93.2894
2.1431 | -93.1256
2.1468 | -92.8672
2.1506 | -92.6168
2.1544 | -92.5263
2.1582 | -92.6416
2.162 | -92.8486
2.1657 | -93.0089
2.1695 | -93.1144
2.1733 | -93.2507
2.1771 | -93.432
2.1809 | -93.6246
2.1846 | -93.7952
2.1884 | -93.9363
2.1922 | -94.0201
2.196 | -94.0347
2.1998 | -93.9949
2.2035 | -93.9416
2.2073 | -93.914
2.2111 | -93.9215
2.2149 | -93.9751
2.2187 | -94.0431
2.2224 | -94.0711
2.2262 | -93.9627
2.23 | -93.7013
2.2338 | -93.3529
2.2376 | -93.0496
2.2413 | -92.8972
2.2451 | -92.9229
2.2489 | -93.0901
2.2527 | -93.3041
2.2564 | -93.4689
2.2602 | -93.5455
2.264 | -93.5555
2.2678 | -93.5385
2.2716 | -93.4907
2.2753 | -93.4319
2.2791 | -93.3625
2.2829 | -93.2607
2.2867 | -93.1309
2.2905 | -93.0084
2.2942 | -92.9373
2.298 | -92.9548
2.3018 | -93.0623
2.3056 | -93.2419
2.3094 | -93.4721
2.3131 | -93.7104
2.3169 | -93.8816
2.3207 | -93.9459
2.3245 | -93.9192
2.3283 | -93.8442
2.332 | -93.7766
2.3358 | -93.7434
2.3396 | -93.737
2.3434 | -93.7866
2.3472 | -93.9116
2.3509 | -94.0772
2.3547 | -94.2224
2.3585 | -94.2881
2.3623 | -94.3039
2.3661 | -94.3443
2.3698 | -94.4419
2.3736 | -94.5709
2.3774 | -94.6855
2.3812 | -94.7584
2.385 | -94.7745
2.3887 | -94.7189
2.3925 | -94.5861
2.3963 | -94.3953
2.4001 | -94.1992
2.4039 | -94.0277
2.4076 | -93.9097
2.4114 | -93.8616
2.4152 | -93.8804
2.419 | -93.9433
2.4228 | -94.0275
2.4265 | -94.1064
2.4303 | -94.1547
2.4341 | -94.1781
2.4379 | -94.1897
2.4417 | -94.2156
2.4454 | -94.2566
2.4492 | -94.2877
2.453 | -94.2896
2.4568 | -94.2641
2.4606 | -94.2275
2.4643 | -94.2039
2.4681 | -94.2083
2.4719 | -94.2169
2.4757 | -94.191
2.4794 | -94.1406
2.4832 | -94.1087
2.487 | -94.1318
2.4908 | -94.1935
2.4946 | -94.2522
2.4983 | -94.2747
2.5021 | -94.2664
2.5059 | -94.2448
2.5097 | -94.2028
2.5135 | -94.1287
2.5172 | -94.0379
2.521 | -93.9451
2.5248 | -93.8786
2.5286 | -93.8642
2.5324 | -93.9111
2.5361 | -94.0248
2.5399 | -94.1944
2.5437 | -94.3845
2.5475 | -94.5271
2.5513 | -94.5781
2.555 | -94.5468
2.5588 | -94.4838
2.5626 | -94.4821
2.5664 | -94.5909
2.5702 | -94.7934
2.5739 | -95.0072
2.5777 | -95.1142
2.5815 | -95.0435
2.5853 | -94.8357
2.5891 | -94.6138
2.5928 | -94.4847
2.5966 | -94.4835
2.6004 | -94.5747
2.6042 | -94.6843
2.608 | -94.7685
2.6117 | -94.7971
2.6155 | -94.7749
2.6193 | -94.7171
2.6231 | -94.6473
2.6269 | -94.5923
2.6306 | -94.5933
2.6344 | -94.6673
2.6382 | -94.7907
2.642 | -94.9184
2.6458 | -94.9819
2.6495 | -94.9605
2.6533 | -94.9014
2.6571 | -94.8287
2.6609 | -94.737
2.6647 | -94.639
2.6684 | -94.5367
2.6722 | -94.4352
2.676 | -94.3514
2.6798 | -94.2902
2.6835 | -94.2631
2.6873 | -94.3063
2.6911 | -94.4068
2.6949 | -94.5261
2.6987 | -94.631
2.7024 | -94.7256
2.7062 | -94.8101
2.71 | -94.8912
2.7138 | -94.9682
2.7176 | -95.0492
2.7213 | -95.1418
2.7251 | -95.1616
2.7289 | -95.167
2.7327 | -95.118
2.7365 | -95.0434
2.7402 | -94.9808
2.744 | -94.9491
2.7478 | -94.9685
2.7516 | -95.0509
2.7554 | -95.1791
2.7591 | -95.3341
2.7629 | -95.4939
2.7667 | -95.6418
2.7705 | -95.759
2.7743 | -95.8344
2.778 | -95.8567
2.7818 | -95.8249
2.7856 | -95.7429
2.7894 | -95.6251
2.7932 | -95.4959
2.7969 | -95.3277
2.8007 | -95.1195
2.8045 | -94.9072
2.8083 | -94.7352
2.8121 | -94.6435
2.8158 | -94.6601
2.8196 | -94.7745
2.8234 | -94.9549
2.8272 | -95.1484
2.831 | -95.3143
2.8347 | -95.4243
2.8385 | -95.4595
2.8423 | -95.4125
2.8461 | -95.3188
2.8499 | -95.2014
2.8536 | -95.0765
2.8574 | -94.9761
2.8612 | -94.9397
2.865 | -94.9687
2.8688 | -95.0557
2.8725 | -95.173
2.8763 | -95.2941
2.8801 | -95.3833
2.8839 | -95.4232
2.8876 | -95.4264
2.8914 | -95.4173
2.8952 | -95.4113
2.899 | -95.4126
2.9028 | -95.4192
2.9065 | -95.4125
2.9103 | -95.3805
2.9141 | -95.3426
2.9179 | -95.3107
2.9217 | -95.3095
2.9254 | -95.3443
2.9292 | -95.4236
2.933 | -95.5214
2.9368 | -95.6027
2.9406 | -95.641
2.9443 | -95.6243
2.9481 | -95.5571
2.9519 | -95.4581
2.9557 | -95.3717
2.9595 | -95.3364
2.9632 | -95.3743
2.967 | -95.4927
2.9708 | -95.6641
2.9746 | -95.862
2.9784 | -96.0461
2.9821 | -96.1893
2.9859 | -96.2677
2.9897 | -96.2869
2.9935 | -96.2783
2.9973 | -96.2931
3.001 | -96.3533
3.0048 | -96.4428
3.0086 | -96.535
3.0124 | -96.6051
3.0162 | -96.6191
3.0199 | -96.5699
3.0237 | -96.4544
3.0275 | -96.2911
3.0313 | -96.1093
3.0351 | -95.9501
3.0388 | -95.8024
3.0426 | -95.6713
3.0464 | -95.5665
3.0502 | -95.5108
3.054 | -95.5187
3.0577 | -95.6002
3.0615 | -95.7153
3.0653 | -95.8335
3.0691 | -95.919
3.0729 | -95.9529
3.0766 | -95.9388
3.0804 | -95.9137
3.0842 | -95.906
3.088 | -95.9635
3.0918 | -96.0762
3.0955 | -96.2131
3.0993 | -96.3348
3.1031 | -96.3912
3.1069 | -96.3468
3.1106 | -96.2473
3.1144 | -96.1212
3.1182 | -96.0284
3.122 | -96.0153
3.1258 | -96.0859
3.1295 | -96.2106
3.1333 | -96.3594
3.1371 | -96.4839
3.1409 | -96.5614
3.1447 | -96.5712
3.1484 | -96.5209
3.1522 | -96.436
3.156 | -96.3437
3.1598 | -96.2716
3.1636 | -96.2665
3.1673 | -96.3223
3.1711 | -96.4154
3.1749 | -96.5121
3.1787 | -96.6037
3.1825 | -96.6959
3.1862 | -96.7859
3.19 | -96.8711
3.1938 | -96.9417
3.1976 | -96.9514
3.2014 | -96.8844
3.2051 | -96.7772
3.2089 | -96.6601
3.2127 | -96.5536
3.2165 | -96.4892
3.2203 | -96.4812
3.224 | -96.5199
3.2278 | -96.5674
3.2316 | -96.599
3.2354 | -96.5992
3.2392 | -96.5511
3.2429 | -96.4846
3.2467 | -96.4568
3.2505 | -96.4944
3.2543 | -96.5906
3.2581 | -96.724
3.2618 | -96.8568
3.2656 | -96.9349
3.2694 | -96.8899
3.2732 | -96.8761
3.277 | -96.8409
3.2807 | -96.7945
3.2845 | -96.744
3.2883 | -96.7049
3.2921 | -96.6657
3.2959 | -96.6204
3.2996 | -96.5629
3.3034 | -96.4975
3.3072 | -96.4358
3.311 | -96.4027
3.3148 | -96.4067
3.3185 | -96.4399
3.3223 | -96.4885
3.3261 | -96.524
3.3299 | -96.5593
3.3336 | -96.5905
3.3374 | -96.6091
3.3412 | -96.618
3.345 | -96.6123
3.3488 | -96.6046
3.3525 | -96.5817
3.3563 | -96.5511
3.3601 | -96.527
3.3639 | -96.5182
3.3677 | -96.5384
3.3714 | -96.5959
3.3752 | -96.6932
3.379 | -96.8073
3.3828 | -96.9221
3.3866 | -97.0115
3.3903 | -97.0747
3.3941 | -97.1277
3.3979 | -97.1636
3.4017 | -97.1881
3.4055 | -97.2073
3.4092 | -97.2362
3.413 | -97.2658
3.4168 | -97.2891
3.4206 | -97.3
3.4244 | -97.2989
3.4281 | -97.2803
3.4319 | -97.257
3.4357 | -97.2064
3.4395 | -97.1301
3.4433 | -97.0264
3.447 | -96.9036
3.4508 | -96.7824
3.4546 | -96.694
3.4584 | -96.6591
3.4622 | -96.6857
3.4659 | -96.7763
3.4697 | -96.9337
3.4735 | -97.1392
3.4773 | -97.3655
3.4811 | -97.5714
3.4848 | -97.7326
3.4886 | -97.8161
3.4924 | -97.8432
3.4962 | -97.8185
3.5 | -97.7644
3.5037 | -97.7042
3.5075 | -97.6437
3.5113 | -97.5881
3.5151 | -97.53
3.5189 | -97.4638
3.5226 | -97.3824
3.5264 | -97.2865
3.5302 | -97.1908
3.534 | -97.1136
3.5377 | -97.062
3.5415 | -97.0396
3.5453 | -97.0431
3.5491 | -97.0546
3.5529 | -97.0726
3.5566 | -97.0953
3.5604 | -97.1209
3.5642 | -97.154
3.568 | -97.1787
3.5718 | -97.1939
3.5755 | -97.1835
3.5793 | -97.141
3.5831 | -97.0942
3.5869 | -97.062
3.5907 | -97.0481
3.5944 | -97.0625
3.5982 | -97.0863
3.602 | -97.1064
3.6058 | -97.1263
3.6096 | -97.1349
3.6133 | -97.1549
3.6171 | -97.173
3.6209 | -97.1731
3.6247 | -97.174
3.6285 | -97.1851
3.6322 | -97.223
3.636 | -97.299
3.6398 | -97.391
3.6436 | -97.4874
3.6474 | -97.5602
3.6511 | -97.5799
3.6549 | -97.5664
3.6587 | -97.5362
3.6625 | -97.5102
3.6663 | -97.4963
3.67 | -97.4869
3.6738 | -97.4913
3.6776 | -97.4956
3.6814 | -97.4762
3.6852 | -97.453
3.6889 | -97.435
3.6927 | -97.4235
3.6965 | -97.4276
3.7003 | -97.462
3.7041 | -97.5285
3.7078 | -97.614
3.7116 | -97.705
3.7154 | -97.7913
3.7192 | -97.8694
3.723 | -97.9411
3.7267 | -97.9959
3.7305 | -98.034
3.7343 | -98.0418
3.7381 | -98.0061
3.7418 | -97.9329
3.7456 | -97.8469
3.7494 | -97.7769
3.7532 | -97.7375
3.757 | -97.7243
3.7607 | -97.7323
3.7645 | -97.7552
3.7683 | -97.7627
3.7721 | -97.7454
3.7759 | -97.6873
3.7796 | -97.598
3.7834 | -97.4921
3.7872 | -97.3914
3.791 | -97.3288
3.7948 | -97.315
3.7985 | -97.3414
3.8023 | -97.4046
3.8061 | -97.4883
3.8099 | -97.5896
3.8137 | -97.6898
3.8174 | -97.789
3.8212 | -97.8853
3.825 | -97.96
3.8288 | -98.0173
3.8326 | -98.0634
3.8363 | -98.1012
3.8401 | -98.1253
3.8439 | -98.1443
3.8477 | -98.1638
3.8515 | -98.1696
3.8552 | -98.1722
3.859 | -98.171
3.8628 | -98.1715
3.8666 | -98.1803
3.8704 | -98.1836
3.8741 | -98.1818
3.8779 | -98.1697
3.8817 | -98.1297
3.8855 | -98.0779
3.8893 | -98.0346
3.893 | -97.996
3.8968 | -97.9718
3.9006 | -97.9768
3.9044 | -97.9978
3.9082 | -98.0208
3.9119 | -98.0491
3.9157 | -98.0978
3.9195 | -98.1613
3.9233 | -98.235
3.9271 | -98.3125
3.9308 | -98.3808
3.9346 | -98.4329
3.9384 | -98.4516
3.9422 | -98.4267
3.946 | -98.3814
3.9497 | -98.3205
3.9535 | -98.2525
3.9573 | -98.2109
3.9611 | -98.2124
3.9648 | -98.262
3.9686 | -98.367
3.9724 | -98.504
3.9762 | -98.6622
3.98 | -98.8141
3.9837 | -98.9276
3.9875 | -99.0057
3.9913 | -99.0278
3.9951 | -98.9947
3.9989 | -98.9162
4.0026 | -98.8159
4.0064 | -98.7085
4.0102 | -98.5915
4.014 | -98.4817
4.0178 | -98.3872
4.0215 | -98.2836
4.0253 | -98.1587
4.0291 | -98.0264
4.0329 | -97.9034
4.0367 | -97.8105
4.0404 | -97.7487
4.0442 | -97.7336
4.048 | -97.7757
4.0518 | -97.8639
4.0556 | -97.9973
4.0593 | -98.1609
4.0631 | -98.3339
4.0669 | -98.4983
4.0707 | -98.6401
4.0745 | -98.7515
4.0782 | -98.8223
4.082 | -98.8523
4.0858 | -98.8583
4.0896 | -98.8482
4.0934 | -98.8227
4.0971 | -98.7931
4.1009 | -98.7673
4.1047 | -98.7489
4.1085 | -98.7228
4.1123 | -98.681
4.116 | -98.6214
4.1198 | -98.5446
4.1236 | -98.4621
4.1274 | -98.3828
4.1312 | -98.3126
4.1349 | -98.2626
4.1387 | -98.2313
4.1425 | -98.2294
4.1463 | -98.27
4.1501 | -98.3506
4.1538 | -98.4658
4.1576 | -98.6111
4.1614 | -98.7723
4.1652 | -98.938
4.1689 | -99.1011
4.1727 | -99.2514
4.1765 | -99.3688
4.1803 | -99.4243
4.1841 | -99.4173
4.1878 | -99.3389
4.1916 | -99.186
4.1954 | -98.9731
4.1992 | -98.736
4.203 | -98.513
4.2067 | -98.3217
4.2105 | -98.1867
4.2143 | -98.1244
4.2181 | -98.135
4.2219 | -98.2097
4.2256 | -98.3336
4.2294 | -98.4827
4.2332 | -98.6354
4.237 | -98.7864
4.2408 | -98.9165
4.2445 | -99.0187
4.2483 | -99.0849
4.2521 | -99.1188
4.2559 | -99.1297
4.2597 | -99.1244
4.2634 | -99.0954
4.2672 | -99.047
4.271 | -98.9877
4.2748 | -98.916
4.2786 | -98.8381
4.2823 | -98.7754
4.2861 | -98.7379
4.2899 | -98.7311
4.2937 | -98.7598
4.2975 | -98.8144
4.3012 | -98.8795
4.305 | -98.9567
4.3088 | -99.0432
4.3126 | -99.1259
4.3164 | -99.1962
4.3201 | -99.2586
4.3239 | -99.2949
4.3277 | -99.2991
4.3315 | -99.2626
4.3353 | -99.1867
4.339 | -99.08
4.3428 | -98.96
4.3466 | -98.8535
4.3504 | -98.7807
4.3542 | -98.7587
4.3579 | -98.7835
4.3617 | -98.8855
4.3655 | -98.9625
4.3693 | -99.0481
4.3731 | -99.126
4.3768 | -99.1954
4.3806 | -99.2598
4.3844 | -99.3154
4.3882 | -99.3554
4.3919 | -99.3939
4.3957 | -99.4339
4.3995 | -99.474
4.4033 | -99.5103
4.4071 | -99.5358
4.4108 | -99.5525
4.4146 | -99.5451
4.4184 | -99.5116
4.4222 | -99.4517
4.426 | -99.3701
4.4297 | -99.2744
4.4335 | -99.1619
4.4373 | -99.0422
4.4411 | -98.9406
4.4449 | -98.8645
4.4486 | -98.8032
4.4524 | -98.7608
4.4562 | -98.7429
4.46 | -98.7472
4.4638 | -98.7664
4.4675 | -98.7951
4.4713 | -98.827
4.4751 | -98.8572
4.4789 | -98.8916
4.4827 | -98.928
4.4864 | -98.9607
4.4902 | -98.9936
4.494 | -99.0124
4.4978 | -99.0142
4.5016 | -98.9969
4.5053 | -98.9654
4.5091 | -98.9268
4.5129 | -98.8847
4.5167 | -98.8524
4.5205 | -98.8447
4.5242 | -98.8696
4.528 | -98.9293
4.5318 | -99.0177
4.5356 | -99.1251
4.5394 | -99.2363
4.5431 | -99.3486
4.5469 | -99.4533
4.5507 | -99.5255
4.5545 | -99.5547
4.5583 | -99.5496
4.562 | -99.5176
4.5658 | -99.4668
4.5696 | -99.4154
4.5734 | -99.3766
4.5772 | -99.37
4.5809 | -99.4001
4.5847 | -99.4576
4.5885 | -99.5429
4.5923 | -99.6389
4.596 | -99.7291
4.5998 | -99.8085
4.6036 | -99.8702
4.6074 | -99.9079
4.6112 | -99.931
4.6149 | -99.9336
4.6187 | -99.9118
4.6225 | -99.8724
4.6263 | -99.8258
4.6301 | -99.7625
4.6338 | -99.6876
4.6376 | -99.611
4.6414 | -99.5302
4.6452 | -99.4557
4.649 | -99.3892
4.6527 | -99.339
4.6565 | -99.3013
4.6603 | -99.267
4.6641 | -99.2463
4.6679 | -99.2357
4.6716 | -99.2403
4.6754 | -99.2711
4.6792 | -99.3277
4.683 | -99.3914
4.6868 | -99.447
4.6905 | -99.5
4.6943 | -99.5436
4.6981 | -99.5648
4.7019 | -99.5751
4.7057 | -99.5734
4.7094 | -99.5722
4.7132 | -99.5694
4.717 | -99.5598
4.7208 | -99.5474
4.7246 | -99.5204
4.7283 | -99.4836
4.7321 | -99.4336
4.7359 | -99.3772
4.7397 | -99.3162
4.7435 | -99.2471
4.7472 | -99.1882
4.751 | -99.1483
4.7548 | -99.1377
4.7586 | -99.1586
4.7624 | -99.2017
4.7661 | -99.2679
4.7699 | -99.3492
4.7737 | -99.4335
4.7775 | -99.5158
4.7813 | -99.5866
4.785 | -99.6397
4.7888 | -99.676
4.7926 | -99.7032
4.7964 | -99.7388
4.8002 | -99.7731
4.8039 | -99.8089
4.8077 | -99.844
4.8115 | -99.8898
4.8153 | -99.9426
4.819 | -99.9774
4.8228 | -99.9907
4.8266 | -99.9839
4.8304 | -99.9587
4.8342 | -99.9143
4.8379 | -99.859
4.8417 | -99.8057
4.8455 | -99.7575
4.8493 | -99.7263
4.8531 | -99.7107
4.8568 | -99.7108
4.8606 | -99.7207
4.8644 | -99.7315
4.8682 | -99.7429
4.872 | -99.7461
4.8757 | -99.7381
4.8795 | -99.7289
4.8833 | -99.7181
4.8871 | -99.7076
4.8909 | -99.7017
4.8946 | -99.7024
4.8984 | -99.6972
4.9022 | -99.6931
4.906 | -99.6985
4.9098 | -99.6957
4.9135 | -99.6946
4.9173 | -99.6942
4.9211 | -99.6957
4.9249 | -99.69
4.9287 | -99.6854
4.9324 | -99.6844
4.9362 | -99.6856
4.94 | -99.6849
4.9438 | -99.6849
4.9476 | -99.6957
4.9513 | -99.7224
4.9551 | -99.7518
4.9589 | -99.7891
4.9627 | -99.8325
4.9665 | -99.8756
4.9702 | -99.9121
4.974 | -99.9384
4.9778 | -99.9454
4.9816 | -99.9423
4.9854 | -99.937
4.9891 | -99.9385
4.9929 | -99.9387
4.9967 | -99.9501
5.0005 | -99.9756
5.0043 | -100.0133
5.008 | -100.0573
5.0118 | -100.1124
5.0156 | -100.1629
5.0194 | -100.2174
5.0232 | -100.2686
5.0269 | -100.2992
5.0307 | -100.2944
5.0345 | -100.2733
5.0383 | -100.2353
5.042 | -100.1767
5.0458 | -100.1016
5.0496 | -100.0251
5.0534 | -99.9574
5.0572 | -99.9028
5.0609 | -99.8717
5.0647 | -99.8691
5.0685 | -99.8923
5.0723 | -99.9378
5.0761 | -99.9989
5.0798 | -100.0594
5.0836 | -100.1221
5.0874 | -100.1791
5.0912 | -100.2166
5.095 | -100.2433
5.0987 | -100.2558
5.1025 | -100.2584
5.1063 | -100.2605
5.1101 | -100.2658
5.1139 | -100.2609
5.1176 | -100.262
5.1214 | -100.2763
5.1252 | -100.2926
5.129 | -100.3053
5.1328 | -100.3226
5.1365 | -100.3428
5.1403 | -100.3626
5.1441 | -100.3851
5.1479 | -100.4041
5.1517 | -100.4198
5.1554 | -100.4155
5.1592 | -100.4008
5.163 | -100.3698
5.1668 | -100.3237
5.1706 | -100.2716
5.1743 | -100.2244
5.1781 | -100.1854
5.1819 | -100.1562
5.1857 | -100.1383
5.1895 | -100.1317
5.1932 | -100.1431
5.197 | -100.1604
5.2008 | -100.1789
5.2046 | -100.1963
5.2084 | -100.207
5.2121 | -100.2155
5.2159 | -100.2245
5.2197 | -100.2359
5.2235 | -100.251
5.2273 | -100.2774
5.231 | -100.3166
5.2348 | -100.3691
5.2386 | -100.4234
5.2424 | -100.4815
5.2461 | -100.5429
5.2499 | -100.6074
5.2537 | -100.6643
5.2575 | -100.7112
5.2613 | -100.7531
5.265 | -100.78
5.2688 | -100.7876
5.2726 | -100.7731
5.2764 | -100.7399
5.2802 | -100.6893
5.2839 | -100.6391
5.2877 | -100.5879
5.2915 | -100.5323
5.2953 | -100.4819
5.2991 | -100.4462
5.3028 | -100.4178
5.3066 | -100.4096
5.3104 | -100.4185
5.3142 | -100.44
5.318 | -100.4738
5.3217 | -100.5093
5.3255 | -100.5495
5.3293 | -100.5804
5.3331 | -100.6028
5.3369 | -100.624
5.3406 | -100.643
5.3444 | -100.6615
5.3482 | -100.681
5.352 | -100.7042
5.3558 | -100.7331
5.3595 | -100.7691
5.3633 | -100.8027
5.3671 | -100.8297
5.3709 | -100.8418
5.3747 | -100.8412
5.3784 | -100.8187
5.3822 | -100.7839
5.386 | -100.7295
5.3898 | -100.6591
5.3936 | -100.5831
5.3973 | -100.5093
5.4011 | -100.4435
5.4049 | -100.3853
5.4087 | -100.3402
5.4125 | -100.3104
5.4162 | -100.3011
5.42 | -100.3115
5.4238 | -100.3397
5.4276 | -100.3821
5.4314 | -100.4344
5.4351 | -100.4937
5.4389 | -100.5442
5.4427 | -100.5715
5.4465 | -100.5958
5.4502 | -100.5697
5.454 | -100.5693
5.4578 | -100.5495
5.4616 | -100.5218
5.4654 | -100.4981
5.4691 | -100.4715
5.4729 | -100.4526
5.4767 | -100.4448
5.4805 | -100.4511
5.4843 | -100.4727
5.488 | -100.5096
5.4918 | -100.5563
5.4956 | -100.6011
5.4994 | -100.6389
5.5032 | -100.676
5.5069 | -100.7178
5.5107 | -100.7547
5.5145 | -100.7832
5.5183 | -100.8065
5.5221 | -100.8275
5.5258 | -100.8495
5.5296 | -100.858
5.5334 | -100.8649
5.5372 | -100.8765
5.541 | -100.8852
5.5447 | -100.9045
5.5485 | -100.9287
5.5523 | -100.9507
5.5561 | -100.9856
5.5599 | -101.0255
5.5636 | -101.0651
5.5674 | -101.1043
5.5712 | -101.1373
5.575 | -101.1624
5.5788 | -101.1837
5.5825 | -101.2058
5.5863 | -101.2089
5.5901 | -101.1899
5.5939 | -101.1537
5.5977 | -101.1088
5.6014 | -101.0545
5.6052 | -100.9938
5.609 | -100.9367
5.6128 | -100.8821
5.6166 | -100.8298
5.6203 | -100.7894
5.6241 | -100.7614
5.6279 | -100.7478
5.6317 | -100.741
5.6355 | -100.746
5.6392 | -100.7607
5.643 | -100.7732
5.6468 | -100.7843
5.6506 | -100.7974
5.6544 | -100.8152
5.6581 | -100.8369
5.6619 | -100.8668
5.6657 | -100.8951
5.6695 | -100.92
5.6733 | -100.9497
5.677 | -100.9747
5.6808 | -100.9952
5.6846 | -101.0146
5.6884 | -101.0237
5.6921 | -101.0217
5.6959 | -101.0083
5.6997 | -100.9931
5.7035 | -100.974
5.7073 | -100.9488
5.711 | -100.9215
5.7148 | -100.8916
5.7186 | -100.8638
5.7224 | -100.8484
5.7262 | -100.8463
5.7299 | -100.8571
5.7337 | -100.8786
5.7375 | -100.9084
5.7413 | -100.9424
5.7451 | -100.9754
5.7488 | -101.0036
5.7526 | -101.0313
5.7564 | -101.0569
5.7602 | -101.0777
5.764 | -101.0903
5.7677 | -101.09
5.7715 | -101.0758
5.7753 | -101.0725
5.7791 | -101.0637
5.7829 | -101.0506
5.7866 | -101.0346
5.7904 | -101.026
5.7942 | -101.0199
5.798 | -101.0211
5.8018 | -101.0387
5.8055 | -101.0689
5.8093 | -101.1049
5.8131 | -101.1461
5.8169 | -101.1946
5.8207 | -101.2483
5.8244 | -101.296
5.8282 | -101.3337
5.832 | -101.3618
5.8358 | -101.3765
5.8396 | -101.3759
5.8433 | -101.3591
5.8471 | -101.3355
5.8509 | -101.3132
5.8547 | -101.2917
5.8585 | -101.279
5.8622 | -101.2627
5.866 | -101.2536
5.8698 | -101.2494
5.8736 | -101.2499
5.8773 | -101.2541
5.8811 | -101.2543
5.8849 | -101.2535
5.8887 | -101.2505
5.8925 | -101.2462
5.8962 | -101.2412
5.9 | -101.2504
5.9038 | -101.269
5.9076 | -101.2991
5.9114 | -101.3432
5.9151 | -101.4027
5.9189 | -101.477
5.9227 | -101.5607
5.9265 | -101.6457
5.9303 | -101.7258
5.934 | -101.7921
5.9378 | -101.8314
5.9416 | -101.8495
5.9454 | -101.8358
5.9492 | -101.7929
5.9529 | -101.7326
5.9567 | -101.6568
5.9605 | -101.5629
5.9643 | -101.4592
5.9681 | -101.3662
5.9718 | -101.273
5.9756 | -101.1938
5.9794 | -101.1227
5.9832 | -101.0614
5.987 | -101.0095
5.9907 | -100.9656
5.9945 | -100.9471
5.9983 | -100.9224
6.0021 | -100.9083
6.0059 | -100.9033
6.0096 | -100.9015
6.0134 | -100.907
6.0172 | -100.9274
6.021 | -100.9628
6.0248 | -101.0132
6.0285 | -101.0801
6.0323 | -101.15
6.0361 | -101.2295
6.0399 | -101.3196
6.0437 | -101.4047
6.0474 | -101.4815
6.0512 | -101.5566
6.055 | -101.6171
6.0588 | -101.6532
6.0626 | -101.667
6.0663 | -101.6539
6.0701 | -101.631
6.0739 | -101.6006
6.0777 | -101.5622
6.0815 | -101.5287
6.0852 | -101.4988
6.089 | -101.4844
6.0928 | -101.4916
6.0966 | -101.5155
6.1003 | -101.552
6.1041 | -101.6007
6.1079 | -101.645
6.1117 | -101.6867
6.1155 | -101.7213
6.1192 | -101.7439
6.123 | -101.7556
6.1268 | -101.7624
6.1306 | -101.7605
6.1344 | -101.7501
6.1381 | -101.7387
6.1419 | -101.7318
6.1457 | -101.7212
6.1495 | -101.7131
6.1533 | -101.7075
6.157 | -101.7062
6.1608 | -101.7047
6.1646 | -101.6963
6.1684 | -101.6853
6.1722 | -101.6663
6.1759 | -101.636
6.1797 | -101.6073
6.1835 | -101.5706
6.1873 | -101.5318
6.1911 | -101.4914
6.1948 | -101.4549
6.1986 | -101.4263
6.2024 | -101.4098
6.2062 | -101.3986
6.21 | -101.3951
6.2137 | -101.3986
6.2175 | -101.4059
6.2213 | -101.4172
6.2251 | -101.4326
6.2289 | -101.4473
6.2326 | -101.4621
6.2364 | -101.4751
6.2402 | -101.4866
6.244 | -101.5
6.2478 | -101.5105
6.2515 | -101.5172
6.2553 | -101.5209
6.2591 | -101.5259
6.2629 | -101.5302
6.2667 | -101.5418
6.2704 | -101.5531
6.2742 | -101.5592
6.278 | -101.569
6.2818 | -101.5809
6.2856 | -101.5958
6.2893 | -101.6104
6.2931 | -101.6179
6.2969 | -101.631
6.3007 | -101.6498
6.3045 | -101.6708
6.3082 | -101.6951
6.312 | -101.7232
6.3158 | -101.7467
6.3196 | -101.7663
6.3233 | -101.7798
6.3271 | -101.7871
6.3309 | -101.7851
6.3347 | -101.7759
6.3385 | -101.766
6.3422 | -101.7513
6.346 | -101.7401
6.3498 | -101.7361
6.3536 | -101.7356
6.3574 | -101.742
6.3611 | -101.7551
6.3649 | -101.7785
6.3687 | -101.8085
6.3725 | -101.8363
6.3763 | -101.8614
6.38 | -101.8909
6.3838 | -101.9242
6.3876 | -101.9579
6.3914 | -101.9829
6.3952 | -102.0081
6.3989 | -102.0302
6.4027 | -102.0429
6.4065 | -102.0525
6.4103 | -102.059
6.4141 | -102.0545
6.4178 | -102.0441
6.4216 | -102.0327
6.4254 | -102.0229
6.4292 | -102.0205
6.433 | -102.0182
6.4367 | -102.0282
6.4405 | -102.0443
6.4443 | -102.0641
6.4481 | -102.0843
6.4519 | -102.1095
6.4556 | -102.1384
6.4594 | -102.1654
6.4632 | -102.1939
6.467 | -102.2216
6.4708 | -102.2498
6.4745 | -102.2716
6.4783 | -102.2833
6.4821 | -102.289
6.4859 | -102.2906
6.4897 | -102.29
6.4934 | -102.2817
6.4972 | -102.2713
6.501 | -102.2505
6.5048 | -102.2323
6.5086 | -102.217
6.5123 | -102.2078
6.5161 | -102.1905
6.5199 | -102.1736
6.5237 | -102.1501
6.5274 | -102.1242
6.5312 | -102.0911
6.535 | -102.0521
6.5388 | -102.0086
6.5426 | -101.9509
6.5463 | -101.9078
6.5501 | -101.8652
6.5539 | -101.821
6.5577 | -101.7856
6.5615 | -101.7579
6.5652 | -101.7361
6.569 | -101.7237
6.5728 | -101.7226
6.5766 | -101.7326
6.5804 | -101.7524
6.5841 | -101.7794
6.5879 | -101.8098
6.5917 | -101.8524
6.5955 | -101.8967
6.5993 | -101.9414
6.603 | -101.9824
6.6068 | -102.0216
6.6106 | -102.0525
6.6144 | -102.0858
6.6182 | -102.1116
6.6219 | -102.1292
6.6257 | -102.1436
6.6295 | -102.1588
6.6333 | -102.1667
6.6371 | -102.1676
6.6408 | -102.1684
6.6446 | -102.1693
6.6484 | -102.165
6.6522 | -102.1627
6.656 | -102.1687
6.6597 | -102.1772
6.6635 | -102.1861
6.6673 | -102.1991
6.6711 | -102.2084
6.6749 | -102.2157
6.6786 | -102.2209
6.6824 | -102.2216
6.6862 | -102.2165
6.69 | -102.2071
6.6938 | -102.1944
6.6975 | -102.183
6.7013 | -102.172
6.7051 | -102.1634
6.7089 | -102.1605
6.7127 | -102.1687
6.7164 | -102.1841
6.7202 | -102.2066
6.724 | -102.242
6.7278 | -102.2768
6.7316 | -102.3067
6.7353 | -102.3367
6.7391 | -102.3655
6.7429 | -102.3873
6.7467 | -102.4021
6.7504 | -102.413
6.7542 | -102.4143
6.758 | -102.4085
6.7618 | -102.3954
6.7656 | -102.3863
6.7693 | -102.3798
6.7731 | -102.3624
6.7769 | -102.3415
6.7807 | -102.3135
6.7845 | -102.2831
6.7882 | -102.2513
6.792 | -102.2186
6.7958 | -102.1829
6.7996 | -102.1455
6.8034 | -102.1134
6.8071 | -102.0904
6.8109 | -102.0729
6.8147 | -102.0585
6.8185 | -102.0593
6.8223 | -102.0772
6.826 | -102.1007
6.8298 | -102.1372
6.8336 | -102.1867
6.8374 | -102.241
6.8412 | -102.2957
6.8449 | -102.351
6.8487 | -102.4059
6.8525 | -102.458
6.8563 | -102.51
6.8601 | -102.5584
6.8638 | -102.5962
6.8676 | -102.6224
6.8714 | -102.6464
6.8752 | -102.6703
6.879 | -102.6856
6.8827 | -102.6988
6.8865 | -102.707
6.8903 | -102.7092
6.8941 | -102.7121
6.8979 | -102.7183
6.9016 | -102.728
6.9054 | -102.7357
6.9092 | -102.7416
6.913 | -102.7499
6.9168 | -102.7601
6.9205 | -102.7656
6.9243 | -102.761
6.9281 | -102.7574
6.9319 | -102.7546
6.9357 | -102.7476
6.9394 | -102.74
6.9432 | -102.7308
6.947 | -102.7177
6.9508 | -102.7026
6.9546 | -102.688
6.9583 | -102.6685
6.9621 | -102.6491
6.9659 | -102.6336
6.9697 | -102.6106
6.9734 | -102.5847
6.9772 | -102.5634
6.981 | -102.5452
6.9848 | -102.533
6.9886 | -102.5256
6.9923 | -102.5295
6.9961 | -102.5423
6.9999 | -102.5612
7.0037 | -102.584
7.0075 | -102.6056
7.0112 | -102.6264
7.015 | -102.6421
7.0188 | -102.6583
7.0226 | -102.6701
7.0264 | -102.6796
7.0301 | -102.6795
7.0339 | -102.6643
7.0377 | -102.6459
7.0415 | -102.6206
7.0453 | -102.591
7.049 | -102.5624
7.0528 | -102.5317
7.0566 | -102.5012
7.0604 | -102.4779
7.0642 | -102.4611
7.0679 | -102.4476
7.0717 | -102.4395
7.0755 | -102.4291
7.0793 | -102.4165
7.0831 | -102.4031
7.0868 | -102.3963
7.0906 | -102.3873
7.0944 | -102.3819
7.0982 | -102.3784
7.102 | -102.3755
7.1057 | -102.3841
7.1095 | -102.4014
7.1133 | -102.4234
7.1171 | -102.4469
7.1209 | -102.4749
7.1246 | -102.5014
7.1284 | -102.5334
7.1322 | -102.5718
7.136 | -102.6082
7.1398 | -102.6477
7.1435 | -102.687
7.1473 | -102.7294
7.1511 | -102.7733
7.1549 | -102.8152
7.1587 | -102.8574
7.1624 | -102.8965
7.1662 | -102.9313
7.17 | -102.9611
7.1738 | -102.9896
7.1775 | -103.0121
7.1813 | -103.0349
7.1851 | -103.0535
7.1889 | -103.0627
7.1927 | -103.0604
7.1964 | -103.0564
7.2002 | -103.0469
7.204 | -103.034
7.2078 | -103.0176
7.2116 | -102.9999
7.2153 | -102.9741
7.2191 | -102.9441
7.2229 | -102.9198
7.2267 | -102.8977
7.2305 | -102.8767
7.2342 | -102.8543
7.238 | -102.8308
7.2418 | -102.8072
7.2456 | -102.7875
7.2494 | -102.7684
7.2531 | -102.7455
7.2569 | -102.7191
7.2607 | -102.6887
7.2645 | -102.6574
7.2683 | -102.6252
7.272 | -102.596
7.2758 | -102.571
7.2796 | -102.5462
7.2834 | -102.5171
7.2872 | -102.4939
7.2909 | -102.4748
7.2947 | -102.4614
7.2985 | -102.4527
7.3023 | -102.4456
7.3061 | -102.4448
7.3098 | -102.4548
7.3136 | -102.4699
7.3174 | -102.4904
7.3212 | -102.5181
7.325 | -102.551
7.3287 | -102.5899
7.3325 | -102.6259
7.3363 | -102.6615
7.3401 | -102.6977
7.3439 | -102.7313
7.3476 | -102.7629
7.3514 | -102.7955
7.3552 | -102.815
7.359 | -102.8259
7.3628 | -102.8364
7.3665 | -102.8397
7.3703 | -102.8369
7.3741 | -102.8321
7.3779 | -102.8245
7.3816 | -102.809
7.3854 | -102.7892
7.3892 | -102.7714
7.393 | -102.7619
7.3968 | -102.7549
7.4005 | -102.7492
7.4043 | -102.7503
7.4081 | -102.7557
7.4119 | -102.7693
7.4157 | -102.7872
7.4194 | -102.8068
7.4232 | -102.8311
7.427 | -102.8597
7.4308 | -102.8911
7.4346 | -102.9283
7.4383 | -102.9644
7.4421 | -103
7.4459 | -103.0266
7.4497 | -103.0468
7.4535 | -103.0679
7.4572 | -103.0756
7.461 | -103.0738
7.4648 | -103.0732
7.4686 | -103.0634
7.4724 | -103.044
7.4761 | -103.0233
7.4799 | -103.0025
7.4837 | -102.9806
7.4875 | -102.9636
7.4913 | -102.955
7.495 | -102.9505
7.4988 | -102.9494
7.5026 | -102.9491
7.5064 | -102.9508
7.5102 | -102.9548
7.5139 | -102.9628
7.5177 | -102.9724
7.5215 | -102.9848
7.5253 | -102.9966
7.5291 | -103.01
7.5328 | -103.026
7.5366 | -103.0456
7.5404 | -103.0604
7.5442 | -103.0717
7.548 | -103.0844
7.5517 | -103.0974
7.5555 | -103.1063
7.5593 | -103.1095
7.5631 | -103.103
7.5669 | -103.0943
7.5706 | -103.0769
7.5744 | -103.0586
7.5782 | -103.0418
7.582 | -103.0267
7.5858 | -103.0152
7.5895 | -103.0031
7.5933 | -102.9931
7.5971 | -102.9843
7.6009 | -102.9753
7.6046 | -102.9694
7.6084 | -102.9677
7.6122 | -102.9668
7.616 | -102.9649
7.6198 | -102.9645
7.6235 | -102.9659
7.6273 | -102.9723
7.6311 | -102.9778
7.6349 | -102.9799
7.6387 | -102.9836
7.6424 | -102.9892
7.6462 | -102.9947
7.65 | -102.9993
7.6538 | -103.005
7.6576 | -103.004
7.6613 | -102.9973
7.6651 | -102.9897
7.6689 | -102.9814
7.6727 | -102.9692
7.6765 | -102.9566
7.6802 | -102.94
7.684 | -102.9201
7.6878 | -102.8971
7.6916 | -102.8738
7.6954 | -102.8536
7.6991 | -102.8352
7.7029 | -102.8219
7.7067 | -102.8109
7.7105 | -102.8083
7.7143 | -102.8119
7.718 | -102.8167
7.7218 | -102.8217
7.7256 | -102.8351
7.7294 | -102.8513
7.7332 | -102.8695
7.7369 | -102.8873
7.7407 | -102.9
7.7445 | -102.9137
7.7483 | -102.9299
7.7521 | -102.9412
7.7558 | -102.9506
7.7596 | -102.9628
7.7634 | -102.9786
7.7672 | -102.9973
7.771 | -103.0176
7.7747 | -103.0429
7.7785 | -103.0709
7.7823 | -103.0969
7.7861 | -103.1229
7.7899 | -103.1474
7.7936 | -103.1696
7.7974 | -103.1901
7.8012 | -103.205
7.805 | -103.2096
7.8087 | -103.2074
7.8125 | -103.2028
7.8163 | -103.1996
7.8201 | -103.1917
7.8239 | -103.1722
7.8276 | -103.1477
7.8314 | -103.1238
7.8352 | -103.1053
7.839 | -103.088
7.8428 | -103.0689
7.8465 | -103.0577
7.8503 | -103.0545
7.8541 | -103.0545
7.8579 | -103.0666
7.8617 | -103.0822
7.8654 | -103.1049
7.8692 | -103.1417
7.873 | -103.1805
7.8768 | -103.2176
7.8806 | -103.2566
7.8843 | -103.2952
7.8881 | -103.3384
7.8919 | -103.3798
7.8957 | -103.4143
7.8995 | -103.4439
7.9032 | -103.473
7.907 | -103.4943
7.9108 | -103.5125
7.9146 | -103.5352
7.9184 | -103.5572
7.9221 | -103.5777
7.9259 | -103.5932
7.9297 | -103.602
7.9335 | -103.6045
7.9373 | -103.6071
7.941 | -103.6131
7.9448 | -103.6095
7.9486 | -103.5971
7.9524 | -103.5814
7.9562 | -103.5632
7.9599 | -103.547
7.9637 | -103.5311
7.9675 | -103.5167
7.9713 | -103.507
7.9751 | -103.4977
7.9788 | -103.4946
7.9826 | -103.4948
7.9864 | -103.493
7.9902 | -103.4989
7.994 | -103.5082
7.9977 | -103.5132
8.0015 | -103.5225
8.0053 | -103.5325
8.0091 | -103.5427
8.0129 | -103.5523
8.0166 | -103.5559
8.0204 | -103.5557
8.0242 | -103.5519
8.028 | -103.5422
8.0317 | -103.5238
8.0355 | -103.5041
8.0393 | -103.4811
8.0431 | -103.4585
8.0469 | -103.4359
8.0506 | -103.417
8.0544 | -103.4041
8.0582 | -103.3936
8.062 | -103.3878
8.0658 | -103.3828
8.0695 | -103.3768
8.0733 | -103.3766
8.0771 | -103.3777
8.0809 | -103.3837
8.0847 | -103.392
8.0884 | -103.401
8.0922 | -103.4098
8.096 | -103.4169
8.0998 | -103.4207
8.1036 | -103.4238
8.1073 | -103.4293
8.1111 | -103.438
8.1149 | -103.4474
8.1187 | -103.4546
8.1225 | -103.4647
8.1262 | -103.4756
8.13 | -103.494
8.1338 | -103.5197
8.1376 | -103.5384
8.1414 | -103.562
8.1451 | -103.5905
8.1489 | -103.6158
8.1527 | -103.6383
8.1565 | -103.658
8.1603 | -103.6735
8.164 | -103.6832
8.1678 | -103.6873
8.1716 | -103.6852
8.1754 | -103.6763
8.1792 | -103.6446
8.1829 | -103.6297
8.1867 | -103.6172
8.1905 | -103.5985
8.1943 | -103.5746
8.1981 | -103.5467
8.2018 | -103.5104
8.2056 | -103.4747
8.2094 | -103.4446
8.2132 | -103.4157
8.217 | -103.3854
8.2207 | -103.3596
8.2245 | -103.3418
8.2283 | -103.3231
8.2321 | -103.3067
8.2358 | -103.2967
8.2396 | -103.2944
8.2434 | -103.2935
8.2472 | -103.2938
8.251 | -103.2948
8.2547 | -103.2981
8.2585 | -103.3034
8.2623 | -103.3098
8.2661 | -103.3146
8.2699 | -103.3219
8.2736 | -103.3291
8.2774 | -103.3399
8.2812 | -103.3461
8.285 | -103.3499
8.2888 | -103.3519
8.2925 | -103.3556
8.2963 | -103.3622
8.3001 | -103.3648
8.3039 | -103.3709
8.3077 | -103.3791
8.3114 | -103.3862
8.3152 | -103.3993
8.319 | -103.4159
8.3228 | -103.438
8.3266 | -103.4672
8.3303 | -103.4916
8.3341 | -103.5125
8.3379 | -103.534
8.3417 | -103.5571
8.3455 | -103.5824
8.3492 | -103.6046
8.353 | -103.6237
8.3568 | -103.6397
8.3606 | -103.6544
8.3644 | -103.6634
8.3681 | -103.6707
8.3719 | -103.6786
8.3757 | -103.6753
8.3795 | -103.6712
8.3833 | -103.6653
8.387 | -103.6502
8.3908 | -103.6247
8.3946 | -103.6025
8.3984 | -103.5854
8.4022 | -103.5643
8.4059 | -103.5448
8.4097 | -103.5262
8.4135 | -103.5083
8.4173 | -103.4925
8.4211 | -103.4803
8.4248 | -103.4762
8.4286 | -103.4679
8.4324 | -103.4585
8.4362 | -103.4561
8.44 | -103.4572
8.4437 | -103.4623
8.4475 | -103.4709
8.4513 | -103.4817
8.4551 | -103.4956
8.4588 | -103.5093
8.4626 | -103.5234
8.4664 | -103.5356
8.4702 | -103.5436
8.474 | -103.555
8.4777 | -103.5649
8.4815 | -103.5736
8.4853 | -103.5772
8.4891 | -103.5811
8.4929 | -103.5863
8.4966 | -103.5932
8.5004 | -103.6039
8.5042 | -103.61
8.508 | -103.6288
8.5118 | -103.6532
8.5155 | -103.6857
8.5193 | -103.7137
8.5231 | -103.7364
8.5269 | -103.7637
8.5307 | -103.7891
8.5344 | -103.8151
8.5382 | -103.8364
8.542 | -103.8523
8.5458 | -103.8617
8.5496 | -103.8665
8.5533 | -103.8717
8.5571 | -103.8699
8.5609 | -103.8636
8.5647 | -103.8549
8.5685 | -103.8387
8.5722 | -103.8207
8.576 | -103.8037
8.5798 | -103.7854
8.5836 | -103.7661
8.5874 | -103.7467
8.5911 | -103.7249
8.5949 | -103.6981
8.5987 | -103.6641
8.6025 | -103.6369
8.6063 | -103.6112
8.61 | -103.5861
8.6138 | -103.5599
8.6176 | -103.5401
8.6214 | -103.5256
8.6252 | -103.5167
8.6289 | -103.5106
8.6327 | -103.5069
8.6365 | -103.5132
8.6403 | -103.5277
8.6441 | -103.5488
8.6478 | -103.5716
8.6516 | -103.593
8.6554 | -103.618
8.6592 | -103.6392
8.663 | -103.659
8.6667 | -103.6734
8.6705 | -103.6814
8.6743 | -103.6842
8.6781 | -103.6813
8.6818 | -103.68
8.6856 | -103.6749
8.6894 | -103.6661
8.6932 | -103.6583
8.697 | -103.6444
8.7007 | -103.6312
8.7045 | -103.6206
8.7083 | -103.6097
8.7121 | -103.5939
8.7159 | -103.5764
8.7196 | -103.5616
8.7234 | -103.5461
8.7272 | -103.5369
8.731 | -103.5258
8.7348 | -103.5104
8.7385 | -103.4968
8.7423 | -103.4893
8.7461 | -103.4836
8.7499 | -103.4852
8.7537 | -103.4903
8.7574 | -103.5032
8.7612 | -103.5206
8.765 | -103.5411
8.7688 | -103.5649
8.7726 | -103.5917
8.7763 | -103.6224
8.7801 | -103.6534
8.7839 | -103.6811
8.7877 | -103.7123
8.7915 | -103.7406
8.7952 | -103.7671
8.799 | -103.7934
8.8028 | -103.8155
8.8066 | -103.8309
8.8104 | -103.8387
8.8141 | -103.8455
8.8179 | -103.8478
8.8217 | -103.8422
8.8255 | -103.8352
8.8293 | -103.8222
8.833 | -103.8114
8.8368 | -103.8029
8.8406 | -103.7988
8.8444 | -103.7995
8.8482 | -103.8006
8.8519 | -103.8061
8.8557 | -103.816
8.8595 | -103.8246
8.8633 | -103.8336
8.8671 | -103.8524
8.8708 | -103.873
8.8746 | -103.8891
8.8784 | -103.9014
8.8822 | -103.9121
8.8859 | -103.926
8.8897 | -103.9384
8.8935 | -103.9432
8.8973 | -103.9484
8.9011 | -103.9577
8.9048 | -103.9683
8.9086 | -103.9726
8.9124 | -103.9725
8.9162 | -103.9724
8.92 | -103.9686
8.9237 | -103.967
8.9275 | -103.9681
8.9313 | -103.9644
8.9351 | -103.9568
8.9389 | -103.9479
8.9426 | -103.9401
8.9464 | -103.9263
8.9502 | -103.9143
8.954 | -103.9077
8.9578 | -103.902
8.9615 | -103.8955
8.9653 | -103.8852
8.9691 | -103.8746
8.9729 | -103.8721
8.9767 | -103.8752
8.9804 | -103.8813
8.9842 | -103.8888
8.988 | -103.9002
8.9918 | -103.9155
8.9956 | -103.9336
8.9993 | -103.9479
9.0031 | -103.9627
9.0069 | -103.9798
9.0107 | -103.9977
9.0145 | -104.0126
9.0182 | -104.0243
9.022 | -104.0329
9.0258 | -104.0424
9.0296 | -104.0458
9.0334 | -104.0487
9.0371 | -104.0523
9.0409 | -104.0575
9.0447 | -104.0605
9.0485 | -104.0577
9.0523 | -104.0478
9.056 | -104.0329
9.0598 | -104.0099
9.0636 | -103.9899
9.0674 | -103.9655
9.0712 | -103.9446
9.0749 | -103.9205
9.0787 | -103.9015
9.0825 | -103.8878
9.0863 | -103.8809
9.09 | -103.8824
9.0938 | -103.8924
9.0976 | -103.9119
9.1014 | -103.9379
9.1052 | -103.9663
9.1089 | -103.9975
9.1127 | -104.034
9.1165 | -104.0738
9.1203 | -104.1124
9.1241 | -104.1515
9.1278 | -104.1902
9.1316 | -104.2259
9.1354 | -104.25
9.1392 | -104.2729
9.143 | -104.2958
9.1467 | -104.3131
9.1505 | -104.3282
9.1543 | -104.3378
9.1581 | -104.3398
9.1619 | -104.335
9.1656 | -104.3145
9.1694 | -104.2857
9.1732 | -104.2545
9.177 | -104.2197
9.1808 | -104.1827
9.1845 | -104.1488
9.1883 | -104.118
9.1921 | -104.0926
9.1959 | -104.0721
9.1997 | -104.059
9.2034 | -104.0482
9.2072 | -104.0443
9.211 | -104.0455
9.2148 | -104.0499
9.2186 | -104.0546
9.2223 | -104.0655
9.2261 | -104.0775
9.2299 | -104.0864
9.2337 | -104.0916
9.2375 | -104.0962
9.2412 | -104.0946
9.245 | -104.086
9.2488 | -104.076
9.2526 | -104.0635
9.2564 | -104.0439
9.2601 | -104.0191
9.2639 | -103.9979
9.2677 | -103.9701
9.2715 | -103.9519
9.2753 | -103.9283
9.279 | -103.9039
9.2828 | -103.8894
9.2866 | -103.8748
9.2904 | -103.8622
9.2942 | -103.8571
9.2979 | -103.8586
9.3017 | -103.8673
9.3055 | -103.8798
9.3093 | -103.8953
9.313 | -103.9134
9.3168 | -103.9341
9.3206 | -103.956
9.3244 | -103.9829
9.3282 | -104.0143
9.3319 | -104.0447
9.3357 | -104.0756
9.3395 | -104.1038
9.3433 | -104.12
9.3471 | -104.137
9.3508 | -104.1551
9.3546 | -104.1693
9.3584 | -104.1765
9.3622 | -104.1828
9.366 | -104.1848
9.3697 | -104.185
9.3735 | -104.1812
9.3773 | -104.1748
9.3811 | -104.1677
9.3849 | -104.1527
9.3886 | -104.1331
9.3924 | -104.1128
9.3962 | -104.0891
9.4 | -104.059
9.4038 | -104.0298
9.4075 | -103.9984
9.4113 | -103.9681
9.4151 | -103.9391
9.4189 | -103.9118
9.4227 | -103.8796
9.4264 | -103.8571
9.4302 | -103.8416
9.434 | -103.8292
9.4378 | -103.8208
9.4416 | -103.8176
9.4453 | -103.8207
9.4491 | -103.8273
9.4529 | -103.8462
9.4567 | -103.8703
9.4605 | -103.8988
9.4642 | -103.9263
9.468 | -103.9557
9.4718 | -103.9885
9.4756 | -104.0234
9.4794 | -104.06
9.4831 | -104.1016
9.4869 | -104.1338
9.4907 | -104.1586
9.4945 | -104.1817
9.4983 | -104.206
9.502 | -104.2341
9.5058 | -104.2628
9.5096 | -104.2863
9.5134 | -104.301
9.5172 | -104.3197
9.5209 | -104.3413
9.5247 | -104.361
9.5285 | -104.3814
9.5323 | -104.4006
9.536 | -104.419
9.5398 | -104.4364
9.5436 | -104.4484
9.5474 | -104.4605
9.5512 | -104.469
9.5549 | -104.4743
9.5587 | -104.4787
9.5625 | -104.487
9.5663 | -104.4875
9.5701 | -104.4867
9.5738 | -104.4869
9.5776 | -104.4877
9.5814 | -104.4854
9.5852 | -104.4832
9.589 | -104.4824
9.5927 | -104.4758
9.5965 | -104.4685
9.6003 | -104.46
9.6041 | -104.4569
9.6079 | -104.4542
9.6116 | -104.4515
9.6154 | -104.4518
9.6192 | -104.4555
9.623 | -104.4599
9.6268 | -104.4691
9.6305 | -104.48
9.6343 | -104.4935
9.6381 | -104.5056
9.6419 | -104.5165
9.6457 | -104.532
9.6494 | -104.5515
9.6532 | -104.5721
9.657 | -104.5902
9.6608 | -104.6036
9.6646 | -104.6226
9.6683 | -104.6406
9.6721 | -104.6531
9.6759 | -104.6706
9.6797 | -104.6891
9.6835 | -104.7074
9.6872 | -104.7214
9.691 | -104.7359
9.6948 | -104.752
9.6986 | -104.766
9.7024 | -104.7742
9.7061 | -104.7883
9.7099 | -104.8024
9.7137 | -104.8099
9.7175 | -104.8201
9.7213 | -104.8312
9.725 | -104.8345
9.7288 | -104.835
9.7326 | -104.8298
9.7364 | -104.8296
9.7401 | -104.8269
9.7439 | -104.8144
9.7477 | -104.7989
9.7515 | -104.7809
9.7553 | -104.7597
9.759 | -104.7334
9.7628 | -104.7088
9.7666 | -104.6735
9.7704 | -104.6358
9.7742 | -104.597
9.7779 | -104.5565
9.7817 | -104.5114
9.7855 | -104.4687
9.7893 | -104.4291
9.7931 | -104.3933
9.7968 | -104.3569
9.8006 | -104.3241
9.8044 | -104.2944
9.8082 | -104.2726
9.812 | -104.2681
9.8157 | -104.2558
9.8195 | -104.2427
9.8233 | -104.2325
9.8271 | -104.2285
9.8309 | -104.2323
9.8346 | -104.2334
9.8384 | -104.2347
9.8422 | -104.2386
9.846 | -104.2478
9.8498 | -104.2546
9.8535 | -104.2601
9.8573 | -104.2713
9.8611 | -104.2798
9.8649 | -104.2898
9.8687 | -104.3056
9.8724 | -104.3184
9.8762 | -104.3304
9.88 | -104.3389
9.8838 | -104.3525
9.8876 | -104.372
9.8913 | -104.3904
9.8951 | -104.4086
9.8989 | -104.4291
9.9027 | -104.4509
9.9065 | -104.4708
9.9102 | -104.4981
9.914 | -104.5275
9.9178 | -104.5555
9.9216 | -104.5784
9.9254 | -104.6024
9.9291 | -104.6219
9.9329 | -104.635
9.9367 | -104.6441
9.9405 | -104.6526
9.9443 | -104.6592
9.948 | -104.6633
9.9518 | -104.6668
9.9556 | -104.6716
9.9594 | -104.679
9.9631 | -104.6886
9.9669 | -104.6991
9.9707 | -104.713
9.9745 | -104.7292
9.9783 | -104.7442
9.982 | -104.7626
9.9858 | -104.7833
9.9896 | -104.8037
9.9934 | -104.8276
9.9972 | -104.8499
10.0009 | -104.8732
10.0047 | -104.8918
10.0085 | -104.9115
10.0123 | -104.9314
10.0161 | -104.9493
10.0198 | -104.9651
10.0236 | -104.9808
10.0274 | -104.9955
10.0312 | -105.0044
10.035 | -105.0108
10.0387 | -105.0192
10.0425 | -105.0284
10.0463 | -105.0309
10.0501 | -105.0358
10.0539 | -105.0432
10.0576 | -105.0454
10.0614 | -105.0445
10.0652 | -105.0449
10.069 | -105.049
10.0728 | -105.0475
10.0765 | -105.0453
10.0803 | -105.0441
10.0841 | -105.0398
10.0879 | -105.0317
10.0917 | -105.0209
10.0954 | -105.0092
10.0992 | -104.9906
10.103 | -104.9698
10.1068 | -104.9502
10.1106 | -104.9291
10.1143 | -104.9097
10.1181 | -104.889
10.1219 | -104.8705
10.1257 | -104.8534
10.1295 | -104.8409
10.1332 | -104.8296
10.137 | -104.8179
10.1408 | -104.8049
10.1446 | -104.7949
10.1484 | -104.7887
10.1521 | -104.7812
10.1559 | -104.7698
10.1597 | -104.7628
10.1635 | -104.7601
10.1672 | -104.7587
10.171 | -104.7593
10.1748 | -104.7555
10.1786 | -104.7513
10.1824 | -104.7436
10.1861 | -104.7331
10.1899 | -104.7221
10.1937 | -104.7141
10.1975 | -104.7025
10.2013 | -104.6908
10.205 | -104.6788
10.2088 | -104.6656
10.2126 | -104.6518
10.2164 | -104.6388
10.2202 | -104.631
10.2239 | -104.6237
10.2277 | -104.6192
10.2315 | -104.6197
10.2353 | -104.6195
10.2391 | -104.6242
10.2428 | -104.635
10.2466 | -104.6542
10.2504 | -104.676
10.2542 | -104.6954
10.258 | -104.7159
10.2617 | -104.7346
10.2655 | -104.7587
10.2693 | -104.7822
10.2731 | -104.8022
10.2769 | -104.8161
10.2806 | -104.8288
10.2844 | -104.8421
10.2882 | -104.8504
10.292 | -104.8555
10.2958 | -104.8566
10.2995 | -104.8574
10.3033 | -104.8572
10.3071 | -104.8477
10.3109 | -104.8375
10.3147 | -104.8271
10.3184 | -104.8117
10.3222 | -104.7967
10.326 | -104.7875
10.3298 | -104.7823
10.3336 | -104.774
10.3373 | -104.7595
10.3411 | -104.7417
10.3449 | -104.7232
10.3487 | -104.7026
10.3525 | -104.6855
10.3562 | -104.6676
10.36 | -104.6438
10.3638 | -104.6262
10.3676 | -104.6092
10.3713 | -104.5933
10.3751 | -104.5799
10.3789 | -104.5673
10.3827 | -104.5535
10.3865 | -104.5463
10.3902 | -104.5405
10.394 | -104.5363
10.3978 | -104.5367
10.4016 | -104.5411
10.4054 | -104.5496
10.4091 | -104.5599
10.4129 | -104.5741
10.4167 | -104.5854
10.4205 | -104.5974
10.4243 | -104.6136
10.428 | -104.6305
10.4318 | -104.6452
10.4356 | -104.6575
10.4394 | -104.6705
10.4432 | -104.6832
10.4469 | -104.6919
10.4507 | -104.6989
10.4545 | -104.7005
10.4583 | -104.7007
10.4621 | -104.6932
10.4658 | -104.6834
10.4696 | -104.6713
10.4734 | -104.6533
10.4772 | -104.6274
10.481 | -104.6019
10.4847 | -104.5792
10.4885 | -104.5571
10.4923 | -104.5337
10.4961 | -104.5132
10.4999 | -104.4925
10.5036 | -104.4722
10.5074 | -104.4593
10.5112 | -104.4513
10.515 | -104.4449
10.5188 | -104.4376
10.5225 | -104.4377
10.5263 | -104.4428
10.5301 | -104.4517
10.5339 | -104.4645
10.5377 | -104.4825
10.5414 | -104.501
10.5452 | -104.5241
10.549 | -104.5497
10.5528 | -104.5764
10.5566 | -104.6117
10.5603 | -104.6501
10.5641 | -104.6898
10.5679 | -104.7311
10.5717 | -104.7738
10.5755 | -104.8182
10.5792 | -104.8629
10.583 | -104.9067
10.5868 | -104.95
10.5906 | -104.9906
10.5943 | -105.029
10.5981 | -105.0626
10.6019 | -105.0937
10.6057 | -105.1198
10.6095 | -105.1448
10.6132 | -105.1625
10.617 | -105.1743
10.6208 | -105.1815
10.6246 | -105.1834
10.6284 | -105.1862
10.6321 | -105.1832
10.6359 | -105.1716
10.6397 | -105.1631
10.6435 | -105.1556
10.6473 | -105.1463
10.651 | -105.1308
10.6548 | -105.1214
10.6586 | -105.1161
10.6624 | -105.1047
10.6662 | -105.0925
10.6699 | -105.0851
10.6737 | -105.081
10.6775 | -105.0771
10.6813 | -105.0706
10.6851 | -105.0676
10.6888 | -105.0661
10.6926 | -105.0673
10.6964 | -105.0692
10.7002 | -105.0698
10.704 | -105.069
10.7077 | -105.0688
10.7115 | -105.0657
10.7153 | -105.0642
10.7191 | -105.0647
10.7229 | -105.0668
10.7266 | -105.0682
10.7304 | -105.0701
10.7342 | -105.0731
10.738 | -105.0797
10.7418 | -105.0857
10.7455 | -105.0922
10.7493 | -105.0959
10.7531 | -105.0979
10.7569 | -105.104
10.7607 | -105.1136
10.7644 | -105.1186
10.7682 | -105.1201
10.772 | -105.1254
10.7758 | -105.1246
10.7796 | -105.1217
10.7833 | -105.1159
10.7871 | -105.107
10.7909 | -105.099
10.7947 | -105.0846
10.7985 | -105.0661
10.8022 | -105.0514
10.806 | -105.042
10.8098 | -105.0243
10.8136 | -105.003
10.8173 | -104.9859
10.8211 | -104.9711
10.8249 | -104.9523
10.8287 | -104.9338
10.8325 | -104.9152
10.8362 | -104.9027
10.84 | -104.891
10.8438 | -104.8835
10.8476 | -104.8771
10.8514 | -104.875
10.8551 | -104.8779
10.8589 | -104.8832
10.8627 | -104.8896
10.8665 | -104.8997
10.8703 | -104.9095
10.874 | -104.9166
10.8778 | -104.9236
10.8816 | -104.9308
10.8854 | -104.9332
10.8892 | -104.9283
10.8929 | -104.9189
10.8967 | -104.9109
10.9005 | -104.9003
10.9043 | -104.8822
10.9081 | -104.8726
10.9118 | -104.8617
10.9156 | -104.8446
10.9194 | -104.8249
10.9232 | -104.8061
10.927 | -104.7921
10.9307 | -104.7761
10.9345 | -104.7631
10.9383 | -104.7498
10.9421 | -104.7398
10.9459 | -104.733
10.9496 | -104.7237
10.9534 | -104.7188
10.9572 | -104.7187
10.961 | -104.7206
10.9648 | -104.7208
10.9685 | -104.7275
10.9723 | -104.7393
10.9761 | -104.7506
10.9799 | -104.7661
10.9837 | -104.7861
10.9874 | -104.8078
10.9912 | -104.8291
10.995 | -104.8524
10.9988 | -104.8776
11.0026 | -104.9033
11.0063 | -104.9292
11.0101 | -104.9539
11.0139 | -104.9782
11.0177 | -105.0024
11.0214 | -105.03
11.0252 | -105.0557
11.029 | -105.0831
11.0328 | -105.1121
11.0366 | -105.1365
11.0403 | -105.1599
11.0441 | -105.1852
11.0479 | -105.2105
11.0517 | -105.2307
11.0555 | -105.2446
11.0592 | -105.2566
11.063 | -105.2641
11.0668 | -105.2696
11.0706 | -105.2699
11.0744 | -105.2671
11.0781 | -105.267
11.0819 | -105.2655
11.0857 | -105.2647
11.0895 | -105.2538
11.0933 | -105.2435
11.097 | -105.2369
11.1008 | -105.2282
11.1046 | -105.2161
11.1084 | -105.2035
11.1122 | -105.1917
11.1159 | -105.1772
11.1197 | -105.165
11.1235 | -105.1481
11.1273 | -105.1333
11.1311 | -105.1183
11.1348 | -105.1031
11.1386 | -105.0904
11.1424 | -105.0759
11.1462 | -105.0624
11.15 | -105.047
11.1537 | -105.0351
11.1575 | -105.0237
11.1613 | -105.0101
11.1651 | -104.997
11.1689 | -104.9887
11.1726 | -104.9849
11.1764 | -104.9802
11.1802 | -104.9761
11.184 | -104.9693
11.1878 | -104.9647
11.1915 | -104.9632
11.1953 | -104.9651
11.1991 | -104.9709
11.2029 | -104.9762
11.2067 | -104.984
11.2104 | -104.9933
11.2142 | -105.0059
11.218 | -105.0211
11.2218 | -105.0399
11.2256 | -105.0622
11.2293 | -105.0842
11.2331 | -105.108
11.2369 | -105.1332
11.2407 | -105.156
11.2444 | -105.173
11.2482 | -105.1905
11.252 | -105.2136
11.2558 | -105.2309
11.2596 | -105.2373
11.2633 | -105.2438
11.2671 | -105.246
11.2709 | -105.2465
11.2747 | -105.2416
11.2785 | -105.2271
11.2822 | -105.2079
11.286 | -105.1876
11.2898 | -105.1621
11.2936 | -105.1304
11.2974 | -105.0984
11.3011 | -105.0675
11.3049 | -105.032
11.3087 | -104.9986
11.3125 | -104.973
11.3163 | -104.9489
11.32 | -104.9261
11.3238 | -104.9101
11.3276 | -104.899
11.3314 | -104.8904
11.3352 | -104.8833
11.3389 | -104.8804
11.3427 | -104.8821
11.3465 | -104.8832
11.3503 | -104.8869
11.3541 | -104.896
11.3578 | -104.9062
11.3616 | -104.9181
11.3654 | -104.9334
11.3692 | -104.95
11.373 | -104.9697
11.3767 | -104.9897
11.3805 | -105.0039
11.3843 | -105.0148
11.3881 | -105.0245
11.3919 | -105.0378
11.3956 | -105.0492
11.3994 | -105.057
11.4032 | -105.0621
11.407 | -105.0691
11.4108 | -105.0705
11.4145 | -105.0686
11.4183 | -105.0661
11.4221 | -105.0675
11.4259 | -105.0651
11.4297 | -105.0602
11.4334 | -105.0535
11.4372 | -105.0476
11.441 | -105.0477
11.4448 | -105.0462
11.4485 | -105.0495
11.4523 | -105.0504
11.4561 | -105.0543
11.4599 | -105.056
11.4637 | -105.0617
11.4674 | -105.0687
11.4712 | -105.0806
11.475 | -105.0912
11.4788 | -105.0997
11.4826 | -105.1121
11.4863 | -105.1262
11.4901 | -105.1396
11.4939 | -105.1548
11.4977 | -105.1702
11.5015 | -105.1851
11.5052 | -105.1964
11.509 | -105.2085
11.5128 | -105.2215
11.5166 | -105.2279
11.5204 | -105.2255
11.5241 | -105.2198
11.5279 | -105.2156
11.5317 | -105.2091
11.5355 | -105.202
11.5393 | -105.1966
11.543 | -105.1939
11.5468 | -105.1961
11.5506 | -105.2002
11.5544 | -105.2028
11.5582 | -105.208
11.5619 | -105.2169
11.5657 | -105.2263
11.5695 | -105.2368
11.5733 | -105.2513
11.5771 | -105.2654
11.5808 | -105.2827
11.5846 | -105.3006
11.5884 | -105.3192
11.5922 | -105.3346
11.596 | -105.3468
11.5997 | -105.3549
11.6035 | -105.3587
11.6073 | -105.3531
11.6111 | -105.3451
11.6149 | -105.3373
11.6186 | -105.3273
11.6224 | -105.309
11.6262 | -105.2891
11.63 | -105.2693
11.6338 | -105.242
11.6375 | -105.2159
11.6413 | -105.1908
11.6451 | -105.1671
11.6489 | -105.1476
11.6526 | -105.1269
11.6564 | -105.1086
11.6602 | -105.0942
11.664 | -105.0813
11.6678 | -105.0703
11.6715 | -105.063
11.6753 | -105.0636
11.6791 | -105.0645
11.6829 | -105.0683
11.6867 | -105.0747
11.6904 | -105.0864
11.6942 | -105.103
11.698 | -105.1179
11.7018 | -105.1323
11.7056 | -105.1462
11.7093 | -105.1616
11.7131 | -105.1747
11.7169 | -105.1863
11.7207 | -105.2
11.7245 | -105.2156
11.7282 | -105.2268
11.732 | -105.2404
11.7358 | -105.2575
11.7396 | -105.2741
11.7434 | -105.2856
11.7471 | -105.297
11.7509 | -105.3022
11.7547 | -105.3092
11.7585 | -105.3149
11.7623 | -105.3275
11.766 | -105.3414
11.7698 | -105.3549
11.7736 | -105.3696
11.7774 | -105.3864
11.7812 | -105.3997
11.7849 | -105.4115
11.7887 | -105.4224
11.7925 | -105.4315
11.7963 | -105.4353
11.8001 | -105.4397
11.8038 | -105.4446
11.8076 | -105.4449
11.8114 | -105.4425
11.8152 | -105.4361
11.819 | -105.4269
11.8227 | -105.4166
11.8265 | -105.4066
11.8303 | -105.3963
11.8341 | -105.3817
11.8379 | -105.3659
11.8416 | -105.3499
11.8454 | -105.3334
11.8492 | -105.3161
11.853 | -105.3001
11.8568 | -105.2866
11.8605 | -105.2713
11.8643 | -105.2571
11.8681 | -105.2497
11.8719 | -105.2464
11.8757 | -105.2464
11.8794 | -105.2486
11.8832 | -105.2548
11.887 | -105.2679
11.8908 | -105.2833
11.8945 | -105.3025
11.8983 | -105.3279
11.9021 | -105.3519
11.9059 | -105.3715
11.9097 | -105.394
11.9134 | -105.4228
11.9172 | -105.4596
11.921 | -105.4963
11.9248 | -105.5271
11.9286 | -105.5544
11.9323 | -105.5762
11.9361 | -105.5941
11.9399 | -105.6115
11.9437 | -105.6253
11.9475 | -105.6341
11.9512 | -105.6396
11.955 | -105.6383
11.9588 | -105.6341
11.9626 | -105.6299
11.9664 | -105.624
11.9701 | -105.614
11.9739 | -105.599
11.9777 | -105.5831
11.9815 | -105.5689
11.9853 | -105.5571
11.989 | -105.5469
11.9928 | -105.5532
11.9966 | -105.5527
12.0004 | -105.5568
12.0042 | -105.561
12.0079 | -105.5659
12.0117 | -105.5778
12.0155 | -105.5954
12.0193 | -105.6121
12.0231 | -105.6321
12.0268 | -105.6516
12.0306 | -105.6737
12.0344 | -105.6989
12.0382 | -105.72
12.042 | -105.7379
12.0457 | -105.7563
12.0495 | -105.773
12.0533 | -105.7892
12.0571 | -105.8039
12.0609 | -105.8123
12.0646 | -105.8168
12.0684 | -105.8195
12.0722 | -105.8195
12.076 | -105.8207
12.0798 | -105.82
12.0835 | -105.8132
12.0873 | -105.8007
12.0911 | -105.7836
12.0949 | -105.766
12.0986 | -105.7509
12.1024 | -105.7379
12.1062 | -105.7232
12.11 | -105.7042
12.1138 | -105.6798
12.1175 | -105.6584
12.1213 | -105.6419
12.1251 | -105.6267
12.1289 | -105.6111
12.1327 | -105.5947
12.1364 | -105.5808
12.1402 | -105.5701
12.144 | -105.5586
12.1478 | -105.5477
12.1516 | -105.5394
12.1553 | -105.5304
12.1591 | -105.5201
12.1629 | -105.5104
12.1667 | -105.5023
12.1705 | -105.4951
12.1742 | -105.4881
12.178 | -105.4832
12.1818 | -105.4803
12.1856 | -105.481
12.1894 | -105.4784
12.1931 | -105.4735
12.1969 | -105.4707
12.2007 | -105.4691
12.2045 | -105.4676
12.2083 | -105.4655
12.212 | -105.4599
12.2158 | -105.4565
12.2196 | -105.455
12.2234 | -105.4504
12.2272 | -105.4449
12.2309 | -105.4378
12.2347 | -105.4306
12.2385 | -105.4257
12.2423 | -105.4185
12.2461 | -105.4078
12.2498 | -105.3988
12.2536 | -105.3867
12.2574 | -105.375
12.2612 | -105.3638
12.265 | -105.3517
12.2687 | -105.3378
12.2725 | -105.3247
12.2763 | -105.3095
12.2801 | -105.2943
12.2839 | -105.2807
12.2876 | -105.271
12.2914 | -105.2635
12.2952 | -105.2542
12.299 | -105.2447
12.3027 | -105.2395
12.3065 | -105.238
12.3103 | -105.2387
12.3141 | -105.2414
12.3179 | -105.2434
12.3216 | -105.2498
12.3254 | -105.2592
12.3292 | -105.2696
12.333 | -105.281
12.3368 | -105.2917
12.3405 | -105.2997
12.3443 | -105.3089
12.3481 | -105.3206
12.3519 | -105.3305
12.3557 | -105.3387
12.3594 | -105.3506
12.3632 | -105.3635
12.367 | -105.3752
12.3708 | -105.3903
12.3746 | -105.4085
12.3783 | -105.4287
12.3821 | -105.4471
12.3859 | -105.4627
12.3897 | -105.4831
12.3935 | -105.5029
12.3972 | -105.521
12.401 | -105.5442
12.4048 | -105.5633
12.4086 | -105.5806
12.4124 | -105.6013
12.4161 | -105.6136
12.4199 | -105.6255
12.4237 | -105.6369
12.4275 | -105.642
12.4313 | -105.6449
12.435 | -105.6475
12.4388 | -105.6459
12.4426 | -105.6421
12.4464 | -105.6321
12.4502 | -105.6206
12.4539 | -105.6081
12.4577 | -105.5891
12.4615 | -105.5712
12.4653 | -105.5505
12.4691 | -105.5273
12.4728 | -105.5043
12.4766 | -105.4823
12.4804 | -105.457
12.4842 | -105.4345
12.488 | -105.4156
12.4917 | -105.3985
12.4955 | -105.3848
12.4993 | -105.3733
12.5031 | -105.3625
12.5069 | -105.3499
12.5106 | -105.3431
12.5144 | -105.3419
12.5182 | -105.3441
12.522 | -105.3448
12.5257 | -105.3461
12.5295 | -105.347
12.5333 | -105.3458
12.5371 | -105.3462
12.5409 | -105.3438
12.5446 | -105.3443
12.5484 | -105.3441
12.5522 | -105.3434
12.556 | -105.3433
12.5598 | -105.34
12.5635 | -105.3403
12.5673 | -105.3455
12.5711 | -105.3483
12.5749 | -105.35
12.5787 | -105.3508
12.5824 | -105.3515
12.5862 | -105.3551
12.59 | -105.3579
12.5938 | -105.3584
12.5976 | -105.3619
12.6013 | -105.3701
12.6051 | -105.3803
12.6089 | -105.3884
12.6127 | -105.3938
12.6165 | -105.3951
12.6202 | -105.3961
12.624 | -105.3974
12.6278 | -105.3967
12.6316 | -105.3931
12.6354 | -105.3909
12.6391 | -105.3916
12.6429 | -105.3879
12.6467 | -105.3848
12.6505 | -105.3802
12.6543 | -105.3707
12.658 | -105.3625
12.6618 | -105.358
12.6656 | -105.3502
12.6694 | -105.3431
12.6732 | -105.3336
12.6769 | -105.3244
12.6807 | -105.3145
12.6845 | -105.3075
12.6883 | -105.3033
12.6921 | -105.2984
12.6958 | -105.2948
12.6996 | -105.2908
12.7034 | -105.2909
12.7072 | -105.2942
12.711 | -105.3035
12.7147 | -105.3182
12.7185 | -105.3308
12.7223 | -105.3458
12.7261 | -105.3615
12.7299 | -105.3779
12.7336 | -105.395
12.7374 | -105.4122
12.7412 | -105.4303
12.745 | -105.4472
12.7487 | -105.4651
12.7525 | -105.4803
12.7563 | -105.4938
12.7601 | -105.506
12.7639 | -105.5164
12.7676 | -105.5226
12.7714 | -105.5281
12.7752 | -105.5341
12.779 | -105.5356
12.7828 | -105.5381
12.7865 | -105.5332
12.7903 | -105.5243
12.7941 | -105.5185
12.7979 | -105.5142
12.8017 | -105.5123
12.8054 | -105.5072
12.8092 | -105.5004
12.813 | -105.4967
12.8168 | -105.4919
12.8206 | -105.4875
12.8243 | -105.4842
12.8281 | -105.4817
12.8319 | -105.48
12.8357 | -105.4806
12.8395 | -105.4822
12.8432 | -105.4841
12.847 | -105.4893
12.8508 | -105.5006
12.8546 | -105.5086
12.8584 | -105.5192
12.8621 | -105.5297
12.8659 | -105.5438
12.8697 | -105.5577
12.8735 | -105.572
12.8773 | -105.584
12.881 | -105.5973
12.8848 | -105.6105
12.8886 | -105.6219
12.8924 | -105.6288
12.8962 | -105.6346
12.8999 | -105.6405
12.9037 | -105.6483
12.9075 | -105.6554
12.9113 | -105.661
12.9151 | -105.6637
12.9188 | -105.6681
12.9226 | -105.6719
12.9264 | -105.6731
12.9302 | -105.6804
12.934 | -105.6848
12.9377 | -105.6868
12.9415 | -105.6898
12.9453 | -105.6981
12.9491 | -105.7069
12.9528 | -105.7158
12.9566 | -105.724
12.9604 | -105.7329
12.9642 | -105.7423
12.968 | -105.7479
12.9717 | -105.7568
12.9755 | -105.7664
12.9793 | -105.7735
12.9831 | -105.7787
12.9869 | -105.7816
12.9906 | -105.7821
12.9944 | -105.7825
12.9982 | -105.7837
13.002 | -105.7822
13.0058 | -105.7759
13.0095 | -105.7675
13.0133 | -105.7558
13.0171 | -105.7448
13.0209 | -105.7329
13.0247 | -105.7207
13.0284 | -105.7055
13.0322 | -105.6877
13.036 | -105.6675
13.0398 | -105.6495
13.0436 | -105.6327
13.0473 | -105.6172
13.0511 | -105.5971
13.0549 | -105.5771
13.0587 | -105.558
13.0625 | -105.5408
13.0662 | -105.5217
13.07 | -105.5039
13.0738 | -105.4884
13.0776 | -105.476
13.0814 | -105.4639
13.0851 | -105.4634
13.0889 | -105.4573
13.0927 | -105.4521
13.0965 | -105.4478
13.1003 | -105.4445
13.104 | -105.4463
13.1078 | -105.4493
13.1116 | -105.4525
13.1154 | -105.4619
13.1192 | -105.4706
13.1229 | -105.4792
13.1267 | -105.4928
13.1305 | -105.5094
13.1343 | -105.5236
13.1381 | -105.5377
13.1418 | -105.553
13.1456 | -105.5679
13.1494 | -105.5802
13.1532 | -105.5904
13.157 | -105.6032
13.1607 | -105.6125
13.1645 | -105.6213
13.1683 | -105.6303
13.1721 | -105.638
13.1758 | -105.6432
13.1796 | -105.6477
13.1834 | -105.6487
13.1872 | -105.6465
13.191 | -105.6456
13.1947 | -105.644
13.1985 | -105.6401
13.2023 | -105.6373
13.2061 | -105.632
13.2099 | -105.6248
13.2136 | -105.6155
13.2174 | -105.6064
13.2212 | -105.5974
13.225 | -105.5885
13.2288 | -105.5798
13.2325 | -105.5701
13.2363 | -105.5619
13.2401 | -105.5578
13.2439 | -105.554
13.2477 | -105.5494
13.2514 | -105.5457
13.2552 | -105.5438
13.259 | -105.5427
13.2628 | -105.5417
13.2666 | -105.5397
13.2703 | -105.5384
13.2741 | -105.5382
13.2779 | -105.5361
13.2817 | -105.5346
13.2855 | -105.5316
13.2892 | -105.528
13.293 | -105.524
13.2968 | -105.5198
13.3006 | -105.5132
13.3044 | -105.5022
13.3081 | -105.4902
13.3119 | -105.4755
13.3157 | -105.4636
13.3195 | -105.4496
13.3233 | -105.4333
13.327 | -105.4173
13.3308 | -105.4023
13.3346 | -105.3849
13.3384 | -105.3683
13.3422 | -105.3531
13.3459 | -105.3384
13.3497 | -105.3247
13.3535 | -105.3113
13.3573 | -105.2987
13.3611 | -105.2864
13.3648 | -105.2736
13.3686 | -105.2628
13.3724 | -105.2553
13.3762 | -105.2493
13.3799 | -105.2465
13.3837 | -105.2471
13.3875 | -105.2462
13.3913 | -105.2488
13.3951 | -105.2546
13.3988 | -105.2609
13.4026 | -105.2678
13.4064 | -105.2744
13.4102 | -105.283
13.414 | -105.2956
13.4177 | -105.3077
13.4215 | -105.32
13.4253 | -105.3324
13.4291 | -105.3453
13.4329 | -105.3549
13.4366 | -105.3634
13.4404 | -105.3735
13.4442 | -105.3811
13.448 | -105.3883
13.4518 | -105.3954
13.4555 | -105.4003
13.4593 | -105.4041
13.4631 | -105.4068
13.4669 | -105.4079
13.4707 | -105.4094
13.4744 | -105.4099
13.4782 | -105.4118
13.482 | -105.4169
13.4858 | -105.4217
13.4896 | -105.4222
13.4933 | -105.4236
13.4971 | -105.4289
13.5009 | -105.4303
13.5047 | -105.433
13.5085 | -105.4344
13.5122 | -105.4381
13.516 | -105.4418
13.5198 | -105.4464
13.5236 | -105.4509
13.5274 | -105.4532
13.5311 | -105.4571
13.5349 | -105.4606
13.5387 | -105.4619
13.5425 | -105.4593
13.5463 | -105.46
13.55 | -105.4636
13.5538 | -105.4682
13.5576 | -105.4755
13.5614 | -105.4769
13.5652 | -105.4789
13.5689 | -105.4845
13.5727 | -105.4883
13.5765 | -105.4932
13.5803 | -105.4999
13.584 | -105.508
13.5878 | -105.5188
13.5916 | -105.5315
13.5954 | -105.5441
13.5992 | -105.5584
13.6029 | -105.569
13.6067 | -105.5787
13.6105 | -105.5924
13.6143 | -105.6072
13.6181 | -105.6202
13.6218 | -105.6337
13.6256 | -105.6498
13.6294 | -105.6636
13.6332 | -105.6771
13.637 | -105.6881
13.6407 | -105.6977
13.6445 | -105.7076
13.6483 | -105.7189
13.6521 | -105.7297
13.6559 | -105.7407
13.6596 | -105.7497
13.6634 | -105.757
13.6672 | -105.7622
13.671 | -105.7674
13.6748 | -105.7725
13.6785 | -105.7767
13.6823 | -105.7815
13.6861 | -105.7879
13.6899 | -105.7943
13.6937 | -105.7996
13.6974 | -105.8055
13.7012 | -105.8134
13.705 | -105.8191
13.7088 | -105.8212
13.7126 | -105.8248
13.7163 | -105.8299
13.7201 | -105.8353
13.7239 | -105.8459
13.7277 | -105.855
13.7315 | -105.8617
13.7352 | -105.8708
13.739 | -105.8791
13.7428 | -105.881
13.7466 | -105.8809
13.7504 | -105.8809
13.7541 | -105.8786
13.7579 | -105.8825
13.7617 | -105.8803
13.7655 | -105.8775
13.7693 | -105.8732
13.773 | -105.8662
13.7768 | -105.8575
13.7806 | -105.847
13.7844 | -105.8359
13.7882 | -105.822
13.7919 | -105.8089
13.7957 | -105.7994
13.7995 | -105.7902
13.8033 | -105.7795
13.807 | -105.7658
13.8108 | -105.7516
13.8146 | -105.7411
13.8184 | -105.7308
13.8222 | -105.7216
13.8259 | -105.7156
13.8297 | -105.7114
13.8335 | -105.7085
13.8373 | -105.7079
13.8411 | -105.7114
13.8448 | -105.7153
13.8486 | -105.7187
13.8524 | -105.7233
13.8562 | -105.7326
13.86 | -105.7448
13.8637 | -105.7583
13.8675 | -105.7749
13.8713 | -105.7898
13.8751 | -105.8045
13.8789 | -105.8181
13.8826 | -105.8332
13.8864 | -105.8497
13.8902 | -105.8644
13.894 | -105.8785
13.8978 | -105.8928
13.9015 | -105.9062
13.9053 | -105.9202
13.9091 | -105.9333
13.9129 | -105.9423
13.9167 | -105.9484
13.9204 | -105.9566
13.9242 | -105.9626
13.928 | -105.9668
13.9318 | -105.9734
13.9356 | -105.9788
13.9393 | -105.9848
13.9431 | -105.9927
13.9469 | -105.9997
13.9507 | -106.0065
13.9545 | -106.0158
13.9582 | -106.0229
13.962 | -106.0282
13.9658 | -106.0295
13.9696 | -106.0341
13.9734 | -106.0367
13.9771 | -106.0434
13.9809 | -106.0479
13.9847 | -106.0505
13.9885 | -106.0552
13.9923 | -106.0591
13.996 | -106.06
13.9998 | -106.0614
14.0036 | -106.063
14.0074 | -106.0657
14.0112 | -106.0702
14.0149 | -106.0732
14.0187 | -106.0731
14.0225 | -106.072
14.0263 | -106.0689
14.03 | -106.0654
14.0338 | -106.0636
14.0376 | -106.0629
14.0414 | -106.0592
14.0452 | -106.0518
14.0489 | -106.0463
14.0527 | -106.0445
14.0565 | -106.0406
14.0603 | -106.0376
14.0641 | -106.0356
14.0678 | -106.0334
14.0716 | -106.0328
14.0754 | -106.0296
14.0792 | -106.0244
14.083 | -106.0191
14.0867 | -106.0135
14.0905 | -106.0105
14.0943 | -106.008
14.0981 | -106.0042
14.1019 | -106.0005
14.1056 | -105.9969
14.1094 | -105.997
14.1132 | -105.9997
14.117 | -106.0016
14.1208 | -106.003
14.1245 | -106.0049
14.1283 | -106.0077
14.1321 | -106.0092
14.1359 | -106.0143
14.1397 | -106.0219
14.1434 | -106.0293
14.1472 | -106.0377
14.151 | -106.0463
14.1548 | -106.0574
14.1586 | -106.0689
14.1623 | -106.0823
14.1661 | -106.0935
14.1699 | -106.104
14.1737 | -106.1117
14.1775 | -106.1231
14.1812 | -106.1316
14.185 | -106.1386
14.1888 | -106.1456
14.1926 | -106.1514
14.1964 | -106.1551
14.2001 | -106.1572
14.2039 | -106.1604
14.2077 | -106.1611
14.2115 | -106.1637
14.2153 | -106.1629
14.219 | -106.1575
14.2228 | -106.1526
14.2266 | -106.1482
14.2304 | -106.1442
14.2341 | -106.1408
14.2379 | -106.1356
14.2417 | -106.1274
14.2455 | -106.1168
14.2493 | -106.1087
14.253 | -106.0953
14.2568 | -106.0774
14.2606 | -106.0592
14.2644 | -106.0402
14.2682 | -106.0173
14.2719 | -105.9952
14.2757 | -105.975
14.2795 | -105.9572
14.2833 | -105.9385
14.2871 | -105.9143
14.2908 | -105.8892
14.2946 | -105.8637
14.2984 | -105.8386
14.3022 | -105.814
14.306 | -105.7877
14.3097 | -105.7631
14.3135 | -105.7376
14.3173 | -105.7145
14.3211 | -105.692
14.3249 | -105.6702
14.3286 | -105.6524
14.3324 | -105.636
14.3362 | -105.6226
14.34 | -105.6085
14.3438 | -105.5977
14.3475 | -105.5904
14.3513 | -105.5848
14.3551 | -105.5824
14.3589 | -105.5799
14.3627 | -105.5814
14.3664 | -105.5893
14.3702 | -105.6002
14.374 | -105.6113
14.3778 | -105.6228
14.3816 | -105.6366
14.3853 | -105.6545
14.3891 | -105.6737
14.3929 | -105.6965
14.3967 | -105.72
14.4005 | -105.7459
14.4042 | -105.7739
14.408 | -105.8015
14.4118 | -105.8277
14.4156 | -105.857
14.4194 | -105.8836
14.4231 | -105.9108
14.4269 | -105.9387
14.4307 | -105.9641
14.4345 | -105.986
14.4383 | -106.0104
14.442 | -106.0347
14.4458 | -106.0538
14.4496 | -106.0706
14.4534 | -106.0897
14.4571 | -106.1068
14.4609 | -106.1248
14.4647 | -106.1391
14.4685 | -106.1519
14.4723 | -106.1626
14.476 | -106.1743
14.4798 | -106.1817
14.4836 | -106.1824
14.4874 | -106.1807
14.4912 | -106.1813
14.4949 | -106.1827
14.4987 | -106.1805
14.5025 | -106.1745
14.5063 | -106.1665
14.5101 | -106.1549
14.5138 | -106.1438
14.5176 | -106.1315
14.5214 | -106.1186
14.5252 | -106.1036
14.529 | -106.0891
14.5327 | -106.0738
14.5365 | -106.0593
14.5403 | -106.0457
14.5441 | -106.0306
14.5479 | -106.0175
14.5516 | -106.0035
14.5554 | -105.9893
14.5592 | -105.9773
14.563 | -105.9676
14.5668 | -105.9603
14.5705 | -105.9514
14.5743 | -105.9418
14.5781 | -105.9347
14.5819 | -105.9246
14.5857 | -105.9165
14.5894 | -105.91
14.5932 | -105.9038
14.597 | -105.8974
14.6008 | -105.8885
14.6046 | -105.8809
14.6083 | -105.8753
14.6121 | -105.8689
14.6159 | -105.8618
14.6197 | -105.858
14.6235 | -105.8539
14.6272 | -105.85
14.631 | -105.8471
14.6348 | -105.8445
14.6386 | -105.8424
14.6424 | -105.8405
14.6461 | -105.8335
14.6499 | -105.8283
14.6537 | -105.8267
14.6575 | -105.8257
14.6612 | -105.8227
14.665 | -105.8218
14.6688 | -105.8199
14.6726 | -105.8196
14.6764 | -105.8194
14.6801 | -105.8194
14.6839 | -105.8177
14.6877 | -105.8174
14.6915 | -105.8195
14.6953 | -105.8229
14.699 | -105.8257
14.7028 | -105.8284
14.7066 | -105.8324
14.7104 | -105.8362
14.7142 | -105.8387
14.7179 | -105.8408
14.7217 | -105.8371
14.7255 | -105.837
14.7293 | -105.8346
14.7331 | -105.8325
14.7368 | -105.831
14.7406 | -105.8301
14.7444 | -105.8284
14.7482 | -105.8228
14.752 | -105.8175
14.7557 | -105.813
14.7595 | -105.8081
14.7633 | -105.8024
14.7671 | -105.7988
14.7709 | -105.7972
14.7746 | -105.7925
14.7784 | -105.7908
14.7822 | -105.7921
14.786 | -105.7946
14.7898 | -105.7988
14.7935 | -105.8108
14.7973 | -105.8233
14.8011 | -105.8358
14.8049 | -105.8482
14.8087 | -105.8629
14.8124 | -105.8821
14.8162 | -105.9038
14.82 | -105.9244
14.8238 | -105.9476
14.8276 | -105.973
14.8313 | -105.996
14.8351 | -106.0199
14.8389 | -106.044
14.8427 | -106.0663
14.8465 | -106.0899
14.8502 | -106.1111
14.854 | -106.1278
14.8578 | -106.1449
14.8616 | -106.1602
14.8653 | -106.1761
14.8691 | -106.1909
14.8729 | -106.2041
14.8767 | -106.2128
14.8805 | -106.2215
14.8842 | -106.2288
14.888 | -106.2301
14.8918 | -106.2295
14.8956 | -106.2302
14.8994 | -106.2266
14.9031 | -106.2202
14.9069 | -106.2088
14.9107 | -106.1978
14.9145 | -106.187
14.9183 | -106.1742
14.922 | -106.1631
14.9258 | -106.152
14.9296 | -106.1408
14.9334 | -106.1252
14.9372 | -106.1083
14.9409 | -106.0931
14.9447 | -106.0778
14.9485 | -106.0631
14.9523 | -106.0508
14.9561 | -106.0397
14.9598 | -106.0295
14.9636 | -106.0216
14.9674 | -106.0147
14.9712 | -106.0087
14.975 | -106.0046
14.9787 | -106.0023
14.9825 | -106
14.9863 | -105.9968
14.9901 | -105.9967
14.9939 | -105.995
14.9976 | -105.9963
15.0014 | -105.9966
15.0052 | -105.9979
15.009 | -105.9995
15.0128 | -106.0017
15.0165 | -106.0063
15.0203 | -106.0085
15.0241 | -106.0126
15.0279 | -106.0178
15.0317 | -106.022
15.0354 | -106.028
15.0392 | -106.0354
15.043 | -106.0432
15.0468 | -106.0532
15.0506 | -106.0632
15.0543 | -106.0722
15.0581 | -106.0821
15.0619 | -106.0933
15.0657 | -106.1077
15.0695 | -106.1202
15.0732 | -106.1275
15.077 | -106.1331
15.0808 | -106.1378
15.0846 | -106.1399
15.0883 | -106.1426
15.0921 | -106.1457
15.0959 | -106.1471
15.0997 | -106.1484
15.1035 | -106.1489
15.1072 | -106.1468
15.111 | -106.1436
15.1148 | -106.1365
15.1186 | -106.1279
15.1224 | -106.1202
15.1261 | -106.1065
15.1299 | -106.0927
15.1337 | -106.0788
15.1375 | -106.0593
15.1413 | -106.0396
15.145 | -106.023
15.1488 | -106.0093
15.1526 | -105.993
15.1564 | -105.9755
15.1602 | -105.9587
15.1639 | -105.9427
15.1677 | -105.9255
15.1715 | -105.9074
15.1753 | -105.8919
15.1791 | -105.877
15.1828 | -105.8638
15.1866 | -105.8521
15.1904 | -105.8389
15.1942 | -105.8282
15.198 | -105.8199
15.2017 | -105.8134
15.2055 | -105.8052
15.2093 | -105.7991
15.2131 | -105.7933
15.2169 | -105.7902
15.2206 | -105.791
15.2244 | -105.7898
15.2282 | -105.7903
15.232 | -105.7932
15.2358 | -105.8006
15.2395 | -105.8078
15.2433 | -105.8133
15.2471 | -105.8216
15.2509 | -105.8315
15.2547 | -105.8402
15.2584 | -105.8504
15.2622 | -105.8626
15.266 | -105.8806
15.2698 | -105.8917
15.2736 | -105.9062
15.2773 | -105.92
15.2811 | -105.9322
15.2849 | -105.9425
15.2887 | -105.9567
15.2925 | -105.9723
15.2962 | -105.9851
15.3 | -105.9984
15.3038 | -106.0111
15.3076 | -106.0221
15.3113 | -106.0314
15.3151 | -106.0403
15.3189 | -106.0489
15.3227 | -106.0547
15.3265 | -106.0585
15.3302 | -106.0628
15.334 | -106.0686
15.3378 | -106.0743
15.3416 | -106.0757
15.3454 | -106.0763
15.3491 | -106.0759
15.3529 | -106.0762
15.3567 | -106.0784
15.3605 | -106.0802
15.3643 | -106.0779
15.368 | -106.0763
15.3718 | -106.0763
15.3756 | -106.0761
15.3794 | -106.0759
15.3832 | -106.074
15.3869 | -106.0729
15.3907 | -106.0735
15.3945 | -106.0735
15.3983 | -106.0745
15.4021 | -106.0775
15.4058 | -106.0827
15.4096 | -106.087
15.4134 | -106.0892
15.4172 | -106.0925
15.421 | -106.0985
15.4247 | -106.1082
15.4285 | -106.1153
15.4323 | -106.1209
15.4361 | -106.1265
15.4399 | -106.1296
15.4436 | -106.1346
15.4474 | -106.1392
15.4512 | -106.1431
15.455 | -106.1471
15.4588 | -106.1517
15.4625 | -106.1561
15.4663 | -106.1578
15.4701 | -106.1598
15.4739 | -106.162
15.4777 | -106.1622
15.4814 | -106.1609
15.4852 | -106.1614
15.489 | -106.163
15.4928 | -106.166
15.4966 | -106.1681
15.5003 | -106.1699
15.5041 | -106.1711
15.5079 | -106.1696
15.5117 | -106.1657
15.5154 | -106.1634
15.5192 | -106.1633
15.523 | -106.1622
15.5268 | -106.1616
15.5306 | -106.1563
15.5343 | -106.1527
15.5381 | -106.1535
15.5419 | -106.1518
15.5457 | -106.1512
15.5495 | -106.1497
15.5532 | -106.1417
15.557 | -106.1356
15.5608 | -106.1296
15.5646 | -106.1246
15.5684 | -106.1192
15.5721 | -106.1096
15.5759 | -106.099
15.5797 | -106.0918
15.5835 | -106.0824
15.5873 | -106.0713
15.591 | -106.0593
15.5948 | -106.0476
15.5986 | -106.0381
15.6024 | -106.0268
15.6062 | -106.016
15.6099 | -106.0082
15.6137 | -106.0021
15.6175 | -105.9961
15.6213 | -105.9903
15.6251 | -105.9825
15.6288 | -105.9752
15.6326 | -105.972
15.6364 | -105.9677
15.6402 | -105.9628
15.644 | -105.958
15.6477 | -105.9548
15.6515 | -105.9532
15.6553 | -105.9534
15.6591 | -105.9556
15.6629 | -105.9582
15.6666 | -105.9603
15.6704 | -105.9645
15.6742 | -105.9669
15.678 | -105.9681
15.6818 | -105.9726
15.6855 | -105.9777
15.6893 | -105.9808
15.6931 | -105.9843
15.6969 | -105.9897
15.7007 | -105.9937
15.7044 | -105.9986
15.7082 | -106.002
15.712 | -106.0072
15.7158 | -106.0113
15.7195 | -106.0129
15.7233 | -106.0171
15.7271 | -106.0211
15.7309 | -106.0262
15.7347 | -106.0291
15.7384 | -106.0311
15.7422 | -106.0304
15.746 | -106.0274
15.7498 | -106.026
15.7536 | -106.0276
15.7573 | -106.0253
15.7611 | -106.0194
15.7649 | -106.0156
15.7687 | -106.0144
15.7725 | -106.0137
15.7762 | -106.0107
15.78 | -106.0059
15.7838 | -106.0018
15.7876 | -105.9988
15.7914 | -105.9968
15.7951 | -105.9969
15.7989 | -105.9963
15.8027 | -105.9919
15.8065 | -105.9896
15.8103 | -105.9889
15.814 | -105.9878
15.8178 | -105.9854
15.8216 | -105.9863
15.8254 | -105.9857
15.8292 | -105.9866
15.8329 | -105.9879
15.8367 | -105.9877
15.8405 | -105.9872
15.8443 | -105.9871
15.8481 | -105.9875
15.8518 | -105.9905
15.8556 | -105.9934
15.8594 | -105.996
15.8632 | -105.9976
15.867 | -105.9979
15.8707 | -105.9986
15.8745 | -105.9989
15.8783 | -106.0005
15.8821 | -106.0008
15.8859 | -106.0003
15.8896 | -106.0021
15.8934 | -106.0058
15.8972 | -106.0083
15.901 | -106.0114
15.9048 | -106.0179
15.9085 | -106.0215
15.9123 | -106.0258
15.9161 | -106.0293
15.9199 | -106.0319
15.9237 | -106.0342
15.9274 | -106.0369
15.9312 | -106.0369
15.935 | -106.0361
15.9388 | -106.0342
15.9425 | -106.0352
15.9463 | -106.0376
15.9501 | -106.0361
15.9539 | -106.0346
15.9577 | -106.0325
15.9614 | -106.0301
15.9652 | -106.0256
15.969 | -106.0205
15.9728 | -106.0177
15.9766 | -106.0144
15.9803 | -106.0105
15.9841 | -106.0073
15.9879 | -106.0055
15.9917 | -106.0051
15.9955 | -106.0032
15.9992 | -106.0003
16.003 | -106.0027
16.0068 | -106.0039
16.0106 | -106.0046
16.0144 | -106.0075
16.0181 | -106.0108
16.0219 | -106.0164
16.0257 | -106.0237
16.0295 | -106.0338
16.0333 | -106.042
16.037 | -106.0512
16.0408 | -106.0606
16.0446 | -106.0717
16.0484 | -106.0846
16.0522 | -106.094
16.0559 | -106.1037
16.0597 | -106.1137
16.0635 | -106.1229
16.0673 | -106.1314
16.0711 | -106.1385
16.0748 | -106.1466
16.0786 | -106.1542
16.0824 | -106.1597
16.0862 | -106.1616
16.09 | -106.1616
16.0937 | -106.1647
16.0975 | -106.1663
16.1013 | -106.1657
16.1051 | -106.1638
16.1089 | -106.1599
16.1126 | -106.1548
16.1164 | -106.1454
16.1202 | -106.1356
16.124 | -106.1242
16.1278 | -106.1099
16.1315 | -106.0979
16.1353 | -106.0849
16.1391 | -106.071
16.1429 | -106.0547
16.1467 | -106.0399
16.1504 | -106.0265
16.1542 | -106.0108
16.158 | -105.9958
16.1618 | -105.9808
16.1655 | -105.9684
16.1693 | -105.9574
16.1731 | -105.9451
16.1769 | -105.9321
16.1807 | -105.9219
16.1844 | -105.913
16.1882 | -105.9072
16.192 | -105.9006
16.1958 | -105.8954
16.1996 | -105.8941
16.2033 | -105.8937
16.2071 | -105.895
16.2109 | -105.8967
16.2147 | -105.9003
16.2185 | -105.9066
16.2222 | -105.9159
16.226 | -105.9261
16.2298 | -105.9378
16.2336 | -105.9491
16.2374 | -105.9609
16.2411 | -105.9746
16.2449 | -105.9884
16.2487 | -106.0038
16.2525 | -106.0205
16.2563 | -106.0369
16.26 | -106.05
16.2638 | -106.0632
16.2676 | -106.0763
16.2714 | -106.0898
16.2752 | -106.1028
16.2789 | -106.1117
16.2827 | -106.1193
16.2865 | -106.1269
16.2903 | -106.134
16.2941 | -106.1426
16.2978 | -106.1501
16.3016 | -106.1553
16.3054 | -106.1594
16.3092 | -106.1612
16.313 | -106.1626
16.3167 | -106.1669
16.3205 | -106.1708
16.3243 | -106.1726
16.3281 | -106.1735
16.3319 | -106.1746
16.3356 | -106.1758
16.3394 | -106.1772
16.3432 | -106.1801
16.347 | -106.1844
16.3508 | -106.187
16.3545 | -106.1927
16.3583 | -106.2043
16.3621 | -106.2113
16.3659 | -106.216
16.3696 | -106.2216
16.3734 | -106.2279
16.3772 | -106.2348
16.381 | -106.2442
16.3848 | -106.2491
16.3885 | -106.2545
16.3923 | -106.259
16.3961 | -106.2663
16.3999 | -106.2727
16.4037 | -106.2765
16.4074 | -106.2778
16.4112 | -106.2818
16.415 | -106.2838
16.4188 | -106.2855
16.4226 | -106.2859
16.4263 | -106.2884
16.4301 | -106.2885
16.4339 | -106.2851
16.4377 | -106.2812
16.4415 | -106.2767
16.4452 | -106.2736
16.449 | -106.2691
16.4528 | -106.2599
16.4566 | -106.2482
16.4604 | -106.2406
16.4641 | -106.2352
16.4679 | -106.229
16.4717 | -106.2234
16.4755 | -106.2182
16.4793 | -106.213
16.483 | -106.2078
16.4868 | -106.2025
16.4906 | -106.1989
16.4944 | -106.1989
16.4982 | -106.2024
16.5019 | -106.2056
16.5057 | -106.2055
16.5095 | -106.2063
16.5133 | -106.2124
16.5171 | -106.2186
16.5208 | -106.2253
16.5246 | -106.2323
16.5284 | -106.2372
16.5322 | -106.2456
16.536 | -106.2558
16.5397 | -106.2616
16.5435 | -106.2703
16.5473 | -106.2807
16.5511 | -106.2904
16.5549 | -106.3014
16.5586 | -106.3141
16.5624 | -106.3267
16.5662 | -106.3358
16.57 | -106.3466
16.5738 | -106.3593
16.5775 | -106.3723
16.5813 | -106.3846
16.5851 | -106.3947
16.5889 | -106.4018
16.5926 | -106.41
16.5964 | -106.4182
16.6002 | -106.4292
16.604 | -106.4407
16.6078 | -106.4498
16.6115 | -106.4594
16.6153 | -106.4676
16.6191 | -106.4725
16.6229 | -106.4794
16.6267 | -106.4879
16.6304 | -106.4949
16.6342 | -106.4982
16.638 | -106.5002
16.6418 | -106.5046
16.6456 | -106.5117
16.6493 | -106.5156
16.6531 | -106.5158
16.6569 | -106.5143
16.6607 | -106.5117
16.6645 | -106.5101
16.6682 | -106.5101
16.672 | -106.509
16.6758 | -106.504
16.6796 | -106.4974
16.6834 | -106.4913
16.6871 | -106.4862
16.6909 | -106.4824
16.6947 | -106.4787
16.6985 | -106.4734
16.7023 | -106.4661
16.706 | -106.4597
16.7098 | -106.4508
16.7136 | -106.4374
16.7174 | -106.4248
16.7212 | -106.4149
16.7249 | -106.4044
16.7287 | -106.3941
16.7325 | -106.3845
16.7363 | -106.3736
16.7401 | -106.3647
16.7438 | -106.3582
16.7476 | -106.3495
16.7514 | -106.3401
16.7552 | -106.3336
16.759 | -106.326
16.7627 | -106.3208
16.7665 | -106.3161
16.7703 | -106.3102
16.7741 | -106.3054
16.7779 | -106.3046
16.7816 | -106.2988
16.7854 | -106.2923
16.7892 | -106.288
16.793 | -106.2823
16.7967 | -106.2777
16.8005 | -106.2759
16.8043 | -106.276
16.8081 | -106.2747
16.8119 | -106.2736
16.8156 | -106.274
16.8194 | -106.2765
16.8232 | -106.2811
16.827 | -106.287
16.8308 | -106.2916
16.8345 | -106.2954
16.8383 | -106.3011
16.8421 | -106.3074
16.8459 | -106.3154
16.8497 | -106.326
16.8534 | -106.3381
16.8572 | -106.3525
16.861 | -106.3664
16.8648 | -106.3779
16.8686 | -106.3927
16.8723 | -106.4076
16.8761 | -106.4237
16.8799 | -106.4396
16.8837 | -106.4551
16.8875 | -106.472
16.8912 | -106.4919
16.895 | -106.5096
16.8988 | -106.5269
16.9026 | -106.5454
16.9064 | -106.5617
16.9101 | -106.5754
16.9139 | -106.5906
16.9177 | -106.6065
16.9215 | -106.6235
16.9253 | -106.6365
16.929 | -106.6458
16.9328 | -106.6541
16.9366 | -106.6626
16.9404 | -106.6683
16.9442 | -106.6739
16.9479 | -106.6806
16.9517 | -106.6896
16.9555 | -106.6997
16.9593 | -106.7071
16.9631 | -106.713
16.9668 | -106.7188
16.9706 | -106.7237
16.9744 | -106.7272
16.9782 | -106.7313
16.982 | -106.7346
16.9857 | -106.7404
16.9895 | -106.7461
16.9933 | -106.7509
16.9971 | -106.757
17.0008 | -106.7645
17.0046 | -106.7719
17.0084 | -106.7791
17.0122 | -106.7857
17.016 | -106.793
17.0197 | -106.8023
17.0235 | -106.8139
17.0273 | -106.8241
17.0311 | -106.8344
17.0349 | -106.8433
17.0386 | -106.8502
17.0424 | -106.8575
17.0462 | -106.8643
17.05 | -106.8711
17.0538 | -106.8769
17.0575 | -106.8819
17.0613 | -106.8845
17.0651 | -106.8846
17.0689 | -106.8847
17.0727 | -106.8838
17.0764 | -106.8821
17.0802 | -106.8779
17.084 | -106.8712
17.0878 | -106.8642
17.0916 | -106.8539
17.0953 | -106.8451
17.0991 | -106.8329
17.1029 | -106.8185
17.1067 | -106.8039
17.1105 | -106.7887
17.1142 | -106.7705
17.118 | -106.7529
17.1218 | -106.7346
17.1256 | -106.7148
17.1294 | -106.6968
17.1331 | -106.6795
17.1369 | -106.6607
17.1407 | -106.6421
17.1445 | -106.6209
17.1483 | -106.6012
17.152 | -106.5866
17.1558 | -106.5729
17.1596 | -106.5588
17.1634 | -106.5479
17.1672 | -106.5355
17.1709 | -106.5268
17.1747 | -106.5226
17.1785 | -106.5186
17.1823 | -106.5156
17.1861 | -106.5148
17.1898 | -106.5163
17.1936 | -106.5184
17.1974 | -106.5208
17.2012 | -106.5233
17.205 | -106.5309
17.2087 | -106.5394
17.2125 | -106.548
17.2163 | -106.5576
17.2201 | -106.5694
17.2238 | -106.5811
17.2276 | -106.5931
17.2314 | -106.6046
17.2352 | -106.6181
17.239 | -106.6306
17.2427 | -106.6388
17.2465 | -106.6481
17.2503 | -106.6576
17.2541 | -106.6677
17.2579 | -106.6786
17.2616 | -106.6901
17.2654 | -106.6968
17.2692 | -106.703
17.273 | -106.7104
17.2768 | -106.7143
17.2805 | -106.719
17.2843 | -106.7232
17.2881 | -106.7252
17.2919 | -106.7255
17.2957 | -106.7252
17.2994 | -106.7252
17.3032 | -106.7245
17.307 | -106.7229
17.3108 | -106.722
17.3146 | -106.7204
17.3183 | -106.7185
17.3221 | -106.7174
17.3259 | -106.7155
17.3297 | -106.712
17.3335 | -106.7058
17.3372 | -106.7006
17.341 | -106.6951
17.3448 | -106.6913
17.3486 | -106.6888
17.3524 | -106.6852
17.3561 | -106.6775
17.3599 | -106.6715
17.3637 | -106.6618
17.3675 | -106.6527
17.3713 | -106.6461
17.375 | -106.6365
17.3788 | -106.6276
17.3826 | -106.6196
17.3864 | -106.6113
17.3902 | -106.6025
17.3939 | -106.5928
17.3977 | -106.5793
17.4015 | -106.567
17.4053 | -106.5547
17.4091 | -106.5417
17.4128 | -106.5305
17.4166 | -106.5195
17.4204 | -106.5049
17.4242 | -106.4891
17.4279 | -106.4724
17.4317 | -106.4542
17.4355 | -106.4343
17.4393 | -106.4136
17.4431 | -106.3934
17.4468 | -106.372
17.4506 | -106.3548
17.4544 | -106.3334
17.4582 | -106.3144
17.462 | -106.2987
17.4657 | -106.2846
17.4695 | -106.2697
17.4733 | -106.2539
17.4771 | -106.2383
17.4809 | -106.2237
17.4846 | -106.2104
17.4884 | -106.1983
17.4922 | -106.1892
17.496 | -106.1809
17.4998 | -106.1731
17.5035 | -106.1676
17.5073 | -106.1639
17.5111 | -106.1607
17.5149 | -106.1606
17.5187 | -106.1613
17.5224 | -106.1631
17.5262 | -106.169
17.53 | -106.1764
17.5338 | -106.1831
17.5376 | -106.1923
17.5413 | -106.2041
17.5451 | -106.2171
17.5489 | -106.231
17.5527 | -106.246
17.5565 | -106.2641
17.5602 | -106.283
17.564 | -106.302
17.5678 | -106.3205
17.5716 | -106.3363
17.5754 | -106.3524
17.5791 | -106.3709
17.5829 | -106.3897
17.5867 | -106.4078
17.5905 | -106.4248
17.5943 | -106.4391
17.598 | -106.4545
17.6018 | -106.4706
17.6056 | -106.4882
17.6094 | -106.5035
17.6132 | -106.5175
17.6169 | -106.5294
17.6207 | -106.5391
17.6245 | -106.5492
17.6283 | -106.5568
17.6321 | -106.5635
17.6358 | -106.5692
17.6396 | -106.5725
17.6434 | -106.575
17.6472 | -106.5763
17.651 | -106.5747
17.6547 | -106.5719
17.6585 | -106.5699
17.6623 | -106.5661
17.6661 | -106.5578
17.6698 | -106.5515
17.6736 | -106.5455
17.6774 | -106.5355
17.6812 | -106.5267
17.685 | -106.518
17.6887 | -106.5102
17.6925 | -106.5021
17.6963 | -106.493
17.7001 | -106.4825
17.7039 | -106.4726
17.7076 | -106.4648
17.7114 | -106.4559
17.7152 | -106.445
17.719 | -106.4333
17.7228 | -106.4239
17.7265 | -106.4179
17.7303 | -106.4126
17.7341 | -106.409
17.7379 | -106.4065
17.7417 | -106.4059
17.7454 | -106.4053
17.7492 | -106.4041
17.753 | -106.4031
17.7568 | -106.4013
17.7606 | -106.4002
17.7643 | -106.3998
17.7681 | -106.3992
17.7719 | -106.4007
17.7757 | -106.4025
17.7795 | -106.4037
17.7832 | -106.4053
17.787 | -106.4076
17.7908 | -106.4114
17.7946 | -106.4152
17.7984 | -106.4156
17.8021 | -106.4174
17.8059 | -106.4204
17.8097 | -106.422
17.8135 | -106.4221
17.8173 | -106.4227
17.821 | -106.4266
17.8248 | -106.4317
17.8286 | -106.4354
17.8324 | -106.439
17.8362 | -106.4428
17.8399 | -106.446
17.8437 | -106.45
17.8475 | -106.4536
17.8513 | -106.4554
17.8551 | -106.4569
17.8588 | -106.4573
17.8626 | -106.4605
17.8664 | -106.4671
17.8702 | -106.4715
17.8739 | -106.4767
17.8777 | -106.4824
17.8815 | -106.487
17.8853 | -106.4911
17.8891 | -106.4962
17.8928 | -106.5032
17.8966 | -106.5111
17.9004 | -106.5199
17.9042 | -106.5271
17.908 | -106.5343
17.9117 | -106.5379
17.9155 | -106.5438
17.9193 | -106.5511
17.9231 | -106.5562
17.9269 | -106.5605
17.9306 | -106.5666
17.9344 | -106.5724
17.9382 | -106.577
17.942 | -106.579
17.9458 | -106.5802
17.9495 | -106.5829
17.9533 | -106.5843
17.9571 | -106.5861
17.9609 | -106.5873
17.9647 | -106.5844
17.9684 | -106.5828
17.9722 | -106.5826
17.976 | -106.5814
17.9798 | -106.5775
17.9836 | -106.5738
17.9873 | -106.5679
17.9911 | -106.5588
17.9949 | -106.5521
17.9987 | -106.5468
18.0025 | -106.5436
18.0062 | -106.5391
18.01 | -106.5334
18.0138 | -106.5303
18.0176 | -106.5276
18.0214 | -106.5238
18.0251 | -106.52
18.0289 | -106.5134
18.0327 | -106.508
18.0365 | -106.504
18.0403 | -106.5002
18.044 | -106.4967
18.0478 | -106.4913
18.0516 | -106.4847
18.0554 | -106.4816
18.0592 | -106.4753
18.0629 | -106.4688
18.0667 | -106.4615
18.0705 | -106.4573
18.0743 | -106.4517
18.0781 | -106.4453
18.0818 | -106.4398
18.0856 | -106.4355
18.0894 | -106.4295
18.0932 | -106.4254
18.0969 | -106.4209
18.1007 | -106.4151
18.1045 | -106.4136
18.1083 | -106.413
18.1121 | -106.4094
18.1158 | -106.407
18.1196 | -106.4049
18.1234 | -106.4015
18.1272 | -106.4005
18.131 | -106.3981
18.1347 | -106.3943
18.1385 | -106.3917
18.1423 | -106.3924
18.1461 | -106.3926
18.1499 | -106.3923
18.1536 | -106.391
18.1574 | -106.3924
18.1612 | -106.3953
18.165 | -106.4
18.1688 | -106.4049
18.1725 | -106.4086
18.1763 | -106.412
18.1801 | -106.415
18.1839 | -106.4192
18.1877 | -106.4241
18.1914 | -106.4257
18.1952 | -106.4284
18.199 | -106.433
18.2028 | -106.4375
18.2066 | -106.4404
18.2103 | -106.445
18.2141 | -106.4498
18.2179 | -106.4509
18.2217 | -106.4548
18.2255 | -106.4606
18.2292 | -106.4665
18.233 | -106.4715
18.2368 | -106.4766
18.2406 | -106.4811
18.2444 | -106.4865
18.2481 | -106.4908
18.2519 | -106.4951
18.2557 | -106.4994
18.2595 | -106.5019
18.2633 | -106.504
18.267 | -106.5081
18.2708 | -106.5119
18.2746 | -106.514
18.2784 | -106.5168
18.2822 | -106.5192
18.2859 | -106.5184
18.2897 | -106.5219
18.2935 | -106.5247
18.2973 | -106.5273
18.301 | -106.5294
18.3048 | -106.5305
18.3086 | -106.531
18.3124 | -106.5358
18.3162 | -106.5392
18.3199 | -106.5447
18.3237 | -106.5495
18.3275 | -106.5562
18.3313 | -106.5604
18.3351 | -106.5639
18.3388 | -106.5655
18.3426 | -106.5681
18.3464 | -106.5722
18.3502 | -106.5766
18.354 | -106.5766
18.3577 | -106.5772
18.3615 | -106.5784
18.3653 | -106.576
18.3691 | -106.5755
18.3729 | -106.5785
18.3766 | -106.5795
18.3804 | -106.5776
18.3842 | -106.5742
18.388 | -106.5724
18.3918 | -106.5713
18.3955 | -106.572
18.3993 | -106.5741
18.4031 | -106.5731
18.4069 | -106.5698
18.4107 | -106.5674
18.4144 | -106.5674
18.4182 | -106.5672
18.422 | -106.5657
18.4258 | -106.5635
18.4296 | -106.5623
18.4333 | -106.562
18.4371 | -106.5629
18.4409 | -106.5654
18.4447 | -106.5673
18.4485 | -106.5671
18.4522 | -106.5692
18.456 | -106.5731
18.4598 | -106.5766
18.4636 | -106.5793
18.4674 | -106.5821
18.4711 | -106.5837
18.4749 | -106.5868
18.4787 | -106.5955
18.4825 | -106.6013
18.4863 | -106.6027
18.49 | -106.6048
18.4938 | -106.6082
18.4976 | -106.6137
18.5014 | -106.6192
18.5052 | -106.6252
18.5089 | -106.6329
18.5127 | -106.6389
18.5165 | -106.644
18.5203 | -106.6518
18.524 | -106.6583
18.5278 | -106.6661
18.5316 | -106.6763
18.5354 | -106.6852
18.5392 | -106.6886
18.5429 | -106.6958
18.5467 | -106.7042
18.5505 | -106.7141
18.5543 | -106.7199
18.5581 | -106.7258
18.5618 | -106.731
18.5656 | -106.7348
18.5694 | -106.7388
18.5732 | -106.7438
18.577 | -106.743
18.5807 | -106.7416
18.5845 | -106.7378
18.5883 | -106.7347
18.5921 | -106.7315
18.5959 | -106.7272
18.5996 | -106.7239
18.6034 | -106.7202
18.6072 | -106.7156
18.611 | -106.7082
18.6148 | -106.7012
18.6185 | -106.6978
18.6223 | -106.6945
18.6261 | -106.6892
18.6299 | -106.6845
18.6337 | -106.6786
18.6374 | -106.6727
18.6412 | -106.6684
18.645 | -106.6671
18.6488 | -106.6677
18.6526 | -106.6673
18.6563 | -106.6673
18.6601 | -106.6675
18.6639 | -106.6685
18.6677 | -106.6721
18.6715 | -106.6765
18.6752 | -106.6805
18.679 | -106.6832
18.6828 | -106.6863
18.6866 | -106.6927
18.6904 | -106.6995
18.6941 | -106.7089
18.6979 | -106.7175
18.7017 | -106.7274
18.7055 | -106.7373
18.7093 | -106.7476
18.713 | -106.7583
18.7168 | -106.7696
18.7206 | -106.7793
18.7244 | -106.7903
18.7281 | -106.8013
18.7319 | -106.8124
18.7357 | -106.8235
18.7395 | -106.8311
18.7433 | -106.8369
18.747 | -106.8454
18.7508 | -106.8552
18.7546 | -106.8622
18.7584 | -106.8661
18.7622 | -106.8685
18.7659 | -106.8719
18.7697 | -106.8748
18.7735 | -106.8787
18.7773 | -106.8831
18.7811 | -106.8847
18.7848 | -106.8837
18.7886 | -106.884
18.7924 | -106.8823
18.7962 | -106.8822
18.8 | -106.881
18.8037 | -106.8785
18.8075 | -106.873
18.8113 | -106.8687
18.8151 | -106.8636
18.8189 | -106.8577
18.8226 | -106.8529
18.8264 | -106.8423
18.8302 | -106.8311
18.834 | -106.8227
18.8378 | -106.817
18.8415 | -106.8082
18.8453 | -106.8018
18.8491 | -106.7964
18.8529 | -106.7909
18.8567 | -106.7845
18.8604 | -106.7748
18.8642 | -106.7687
18.868 | -106.763
18.8718 | -106.7596
18.8756 | -106.7546
18.8793 | -106.7511
18.8831 | -106.7492
18.8869 | -106.7476
18.8907 | -106.744
18.8945 | -106.7425
18.8982 | -106.7446
18.902 | -106.7461
18.9058 | -106.7474
18.9096 | -106.7506
18.9134 | -106.7538
18.9171 | -106.7566
18.9209 | -106.7601
18.9247 | -106.7637
18.9285 | -106.7676
18.9322 | -106.7716
18.936 | -106.7756
18.9398 | -106.78
18.9436 | -106.7856
18.9474 | -106.7925
18.9511 | -106.798
18.9549 | -106.8018
18.9587 | -106.8064
18.9625 | -106.8116
18.9663 | -106.8173
18.97 | -106.8214
18.9738 | -106.8255
18.9776 | -106.8295
18.9814 | -106.8347
18.9852 | -106.8375
18.9889 | -106.8401
18.9927 | -106.843
18.9965 | -106.8479
19.0003 | -106.8521
19.0041 | -106.8544
19.0078 | -106.8562
19.0116 | -106.8549
19.0154 | -106.8532
19.0192 | -106.8513
19.023 | -106.8499
19.0267 | -106.8498
19.0305 | -106.8485
19.0343 | -106.8438
19.0381 | -106.84
19.0419 | -106.834
19.0456 | -106.8277
19.0494 | -106.8236
19.0532 | -106.8198
19.057 | -106.816
19.0608 | -106.812
19.0645 | -106.8056
19.0683 | -106.8015
19.0721 | -106.7987
19.0759 | -106.7965
19.0797 | -106.7908
19.0834 | -106.7904
19.0872 | -106.7885
19.091 | -106.7867
19.0948 | -106.7839
19.0986 | -106.7836
19.1023 | -106.7804
19.1061 | -106.7775
19.1099 | -106.7742
19.1137 | -106.7702
19.1175 | -106.7652
19.1212 | -106.7607
19.125 | -106.7582
19.1288 | -106.7553
19.1326 | -106.7526
19.1364 | -106.7495
19.1401 | -106.7453
19.1439 | -106.7395
19.1477 | -106.736
19.1515 | -106.7304
19.1552 | -106.7244
19.159 | -106.7185
19.1628 | -106.7143
19.1666 | -106.7102
19.1704 | -106.7035
19.1741 | -106.6973
19.1779 | -106.6916
19.1817 | -106.6837
19.1855 | -106.6757
19.1893 | -106.6677
19.193 | -106.6598
19.1968 | -106.6529
19.2006 | -106.6462
19.2044 | -106.6361
19.2082 | -106.6245
19.2119 | -106.6157
19.2157 | -106.6078
19.2195 | -106.5981
19.2233 | -106.59
19.2271 | -106.584
19.2308 | -106.5744
19.2346 | -106.5641
19.2384 | -106.556
19.2422 | -106.5479
19.246 | -106.5405
19.2497 | -106.5337
19.2535 | -106.5276
19.2573 | -106.5234
19.2611 | -106.5198
19.2649 | -106.5156
19.2686 | -106.5128
19.2724 | -106.5126
19.2762 | -106.5134
19.28 | -106.5159
19.2838 | -106.5206
19.2875 | -106.525
19.2913 | -106.529
19.2951 | -106.5344
19.2989 | -106.5384
19.3027 | -106.5428
19.3064 | -106.5494
19.3102 | -106.5544
19.314 | -106.56
19.3178 | -106.5667
19.3216 | -106.5722
19.3253 | -106.5751
19.3291 | -106.5779
19.3329 | -106.581
19.3367 | -106.5843
19.3405 | -106.5866
19.3442 | -106.5892
19.348 | -106.5918
19.3518 | -106.5949
19.3556 | -106.5961
19.3593 | -106.5947
19.3631 | -106.5948
19.3669 | -106.5974
19.3707 | -106.5979
19.3745 | -106.5968
19.3782 | -106.5979
19.382 | -106.5958
19.3858 | -106.5913
19.3896 | -106.5896
19.3934 | -106.5874
19.3971 | -106.5838
19.4009 | -106.5787
19.4047 | -106.5745
19.4085 | -106.5718
19.4123 | -106.5667
19.416 | -106.5635
19.4198 | -106.5612
19.4236 | -106.5575
19.4274 | -106.553
19.4312 | -106.5497
19.4349 | -106.5452
19.4387 | -106.5415
19.4425 | -106.5388
19.4463 | -106.5359
19.4501 | -106.5332
19.4538 | -106.5322
19.4576 | -106.5321
19.4614 | -106.5313
19.4652 | -106.5304
19.469 | -106.5305
19.4727 | -106.527
19.4765 | -106.5229
19.4803 | -106.5237
19.4841 | -106.5252
19.4879 | -106.5241
19.4916 | -106.5215
19.4954 | -106.5194
19.4992 | -106.5159
19.503 | -106.5127
19.5068 | -106.5126
19.5105 | -106.5123
19.5143 | -106.511
19.5181 | -106.5089
19.5219 | -106.5058
19.5257 | -106.5026
19.5294 | -106.4974
19.5332 | -106.4927
19.537 | -106.4887
19.5408 | -106.4841
19.5446 | -106.4801
19.5483 | -106.4752
19.5521 | -106.4697
19.5559 | -106.4606
19.5597 | -106.4503
19.5635 | -106.4395
19.5672 | -106.431
19.571 | -106.4221
19.5748 | -106.4144
19.5786 | -106.4061
19.5823 | -106.3957
19.5861 | -106.3841
19.5899 | -106.3725
19.5937 | -106.3614
19.5975 | -106.3506
19.6012 | -106.3407
19.605 | -106.3308
19.6088 | -106.3197
19.6126 | -106.3074
19.6164 | -106.2971
19.6201 | -106.2849
19.6239 | -106.2724
19.6277 | -106.2656
19.6315 | -106.2547
19.6353 | -106.2463
19.639 | -106.2371
19.6428 | -106.228
19.6466 | -106.2192
19.6504 | -106.211
19.6542 | -106.2038
19.6579 | -106.1972
19.6617 | -106.194
19.6655 | -106.1925
19.6693 | -106.1903
19.6731 | -106.188
19.6768 | -106.1873
19.6806 | -106.1891
19.6844 | -106.192
19.6882 | -106.1934
19.692 | -106.1968
19.6957 | -106.2014
19.6995 | -106.2073
19.7033 | -106.214
19.7071 | -106.223
19.7109 | -106.2306
19.7146 | -106.2391
19.7184 | -106.2481
19.7222 | -106.2562
19.726 | -106.2655
19.7298 | -106.2762
19.7335 | -106.2865
19.7373 | -106.297
19.7411 | -106.3085
19.7449 | -106.3214
19.7487 | -106.3337
19.7524 | -106.3479
19.7562 | -106.36
19.76 | -106.3681
19.7638 | -106.3778
19.7676 | -106.3879
19.7713 | -106.3978
19.7751 | -106.4081
19.7789 | -106.4167
19.7827 | -106.4207
19.7864 | -106.4244
19.7902 | -106.4298
19.794 | -106.435
19.7978 | -106.4415
19.8016 | -106.4468
19.8053 | -106.4494
19.8091 | -106.4525
19.8129 | -106.455
19.8167 | -106.4547
19.8205 | -106.453
19.8242 | -106.4499
19.828 | -106.4476
19.8318 | -106.445
19.8356 | -106.4433
19.8394 | -106.4406
19.8431 | -106.4358
19.8469 | -106.4265
19.8507 | -106.418
19.8545 | -106.4119
19.8583 | -106.4059
19.862 | -106.3994
19.8658 | -106.3932
19.8696 | -106.3872
19.8734 | -106.3803
19.8772 | -106.3741
19.8809 | -106.3686
19.8847 | -106.3622
19.8885 | -106.3575
19.8923 | -106.3547
19.8961 | -106.3514
19.8998 | -106.3495
19.9036 | -106.3463
19.9074 | -106.3453
19.9112 | -106.3444
19.915 | -106.3435
19.9187 | -106.3405
19.9225 | -106.3401
19.9263 | -106.3403
19.9301 | -106.3403
19.9339 | -106.3446
19.9376 | -106.3502
19.9414 | -106.3548
19.9452 | -106.3572
19.949 | -106.3586
19.9528 | -106.3633
19.9565 | -106.3692
19.9603 | -106.374
19.9641 | -106.3801
19.9679 | -106.3865
19.9717 | -106.3953
19.9754 | -106.4032
19.9792 | -106.4092
19.983 | -106.4136
19.9868 | -106.4205
19.9906 | -106.4281
19.9943 | -106.4346
19.9981 | -106.4406
20.0019 | -106.4459
20.0057 | -106.4506
20.0094 | -106.4548
20.0132 | -106.457
20.017 | -106.4587
20.0208 | -106.4614
20.0246 | -106.4638
20.0283 | -106.465
20.0321 | -106.4663
20.0359 | -106.4698
20.0397 | -106.4732
20.0435 | -106.4768
20.0472 | -106.4765
20.051 | -106.4756
20.0548 | -106.4735
20.0586 | -106.4713
20.0624 | -106.4693
20.0661 | -106.4696
20.0699 | -106.47
20.0737 | -106.4664
20.0775 | -106.4625
20.0813 | -106.4599
20.085 | -106.4571
20.0888 | -106.4572
20.0926 | -106.4564
20.0964 | -106.454
20.1002 | -106.4534
20.1039 | -106.4541
20.1077 | -106.455
20.1115 | -106.4561
20.1153 | -106.4581
20.1191 | -106.4607
20.1228 | -106.4622
20.1266 | -106.4628
20.1304 | -106.4648
20.1342 | -106.4684
20.138 | -106.469
20.1417 | -106.4675
20.1455 | -106.4685
20.1493 | -106.4716
20.1531 | -106.474
20.1569 | -106.4768
20.1606 | -106.4802
20.1644 | -106.4817
20.1682 | -106.4828
20.172 | -106.4839
20.1758 | -106.4856
20.1795 | -106.4867
20.1833 | -106.4887
20.1871 | -106.4918
20.1909 | -106.4955
20.1947 | -106.4976
20.1984 | -106.5013
20.2022 | -106.5057
20.206 | -106.5121
20.2098 | -106.5175
20.2136 | -106.5217
20.2173 | -106.5251
20.2211 | -106.5248
20.2249 | -106.5268
20.2287 | -106.5294
20.2324 | -106.5319
20.2362 | -106.5354
20.24 | -106.5369
20.2438 | -106.5381
20.2476 | -106.5396
20.2513 | -106.5381
20.2551 | -106.5379
20.2589 | -106.5385
20.2627 | -106.538
20.2665 | -106.5352
20.2702 | -106.5342
20.274 | -106.5327
20.2778 | -106.5312
20.2816 | -106.5297
20.2854 | -106.5256
20.2891 | -106.5219
20.2929 | -106.5192
20.2967 | -106.5154
20.3005 | -106.5095
20.3043 | -106.5032
20.308 | -106.4953
20.3118 | -106.4857
20.3156 | -106.4781
20.3194 | -106.4723
20.3232 | -106.466
20.3269 | -106.4615
20.3307 | -106.4575
20.3345 | -106.4522
20.3383 | -106.4458
20.3421 | -106.4401
20.3458 | -106.4347
20.3496 | -106.4325
20.3534 | -106.4288
20.3572 | -106.4217
20.361 | -106.4164
20.3647 | -106.4104
20.3685 | -106.4045
20.3723 | -106.3994
20.3761 | -106.3921
20.3799 | -106.3862
20.3836 | -106.3803
20.3874 | -106.373
20.3912 | -106.3662
20.395 | -106.3599
20.3988 | -106.3525
20.4025 | -106.3463
20.4063 | -106.3418
20.4101 | -106.3371
20.4139 | -106.3328
20.4177 | -106.3276
20.4214 | -106.3208
20.4252 | -106.3168
20.429 | -106.3153
20.4328 | -106.3127
20.4365 | -106.3083
20.4403 | -106.3032
20.4441 | -106.2982
20.4479 | -106.2964
20.4517 | -106.2972
20.4554 | -106.298
20.4592 | -106.3004
20.463 | -106.3022
20.4668 | -106.3041
20.4706 | -106.3066
20.4743 | -106.3097
20.4781 | -106.3122
20.4819 | -106.3145
20.4857 | -106.3179
20.4895 | -106.3213
20.4932 | -106.3263
20.497 | -106.3311
20.5008 | -106.3358
20.5046 | -106.3414
20.5084 | -106.3483
20.5121 | -106.3563
20.5159 | -106.3624
20.5197 | -106.3689
20.5235 | -106.3753
20.5273 | -106.379
20.531 | -106.3828
20.5348 | -106.3879
20.5386 | -106.3929
20.5424 | -106.3995
20.5462 | -106.4041
20.5499 | -106.4073
20.5537 | -106.4111
20.5575 | -106.4158
20.5613 | -106.4201
20.5651 | -106.4245
20.5688 | -106.4275
20.5726 | -106.4295
20.5764 | -106.4337
20.5802 | -106.4371
20.584 | -106.4387
20.5877 | -106.4391
20.5915 | -106.4399
20.5953 | -106.441
20.5991 | -106.4414
20.6029 | -106.4407
20.6066 | -106.439
20.6104 | -106.434
20.6142 | -106.4275
20.618 | -106.4197
20.6218 | -106.4108
20.6255 | -106.4049
20.6293 | -106.4008
20.6331 | -106.3969
20.6369 | -106.3923
20.6407 | -106.3858
20.6444 | -106.3797
20.6482 | -106.3736
20.652 | -106.3666
20.6558 | -106.3615
20.6595 | -106.3549
20.6633 | -106.3474
20.6671 | -106.3402
20.6709 | -106.3323
20.6747 | -106.3239
20.6784 | -106.3189
20.6822 | -106.315
20.686 | -106.3099
20.6898 | -106.3048
20.6936 | -106.2982
20.6973 | -106.2935
20.7011 | -106.29
20.7049 | -106.286
20.7087 | -106.2837
20.7125 | -106.2816
20.7162 | -106.2784
20.72 | -106.2738
20.7238 | -106.2691
20.7276 | -106.2647
20.7314 | -106.2616
20.7351 | -106.2596
20.7389 | -106.2565
20.7427 | -106.2515
20.7465 | -106.2468
20.7503 | -106.2444
20.754 | -106.2399
20.7578 | -106.2354
20.7616 | -106.2309
20.7654 | -106.2272
20.7692 | -106.2227
20.7729 | -106.2184
20.7767 | -106.2138
20.7805 | -106.2092
20.7843 | -106.2055
20.7881 | -106.2022
20.7918 | -106.1996
20.7956 | -106.1957
20.7994 | -106.1912
20.8032 | -106.1878
20.807 | -106.1851
20.8107 | -106.1807
20.8145 | -106.1776
20.8183 | -106.1762
20.8221 | -106.1756
20.8259 | -106.1757
20.8296 | -106.1764
20.8334 | -106.1768
20.8372 | -106.1745
20.841 | -106.1718
20.8448 | -106.1713
20.8485 | -106.1725
20.8523 | -106.1737
20.8561 | -106.1747
20.8599 | -106.1749
20.8636 | -106.1759
20.8674 | -106.1789
20.8712 | -106.1816
20.875 | -106.1827
20.8788 | -106.186
20.8825 | -106.1912
20.8863 | -106.198
20.8901 | -106.2015
20.8939 | -106.2053
20.8977 | -106.2108
20.9014 | -106.2151
20.9052 | -106.221
20.909 | -106.2275
20.9128 | -106.236
20.9166 | -106.2455
20.9203 | -106.2542
20.9241 | -106.2628
20.9279 | -106.2706
20.9317 | -106.278
20.9355 | -106.2851
20.9392 | -106.2934
20.943 | -106.3026
20.9468 | -106.3131
20.9506 | -106.3201
20.9544 | -106.3259
20.9581 | -106.3342
20.9619 | -106.3427
20.9657 | -106.3508
20.9695 | -106.3595
20.9733 | -106.3642
20.977 | -106.3671
20.9808 | -106.3708
20.9846 | -106.3753
20.9884 | -106.3812
20.9922 | -106.3831
20.9959 | -106.3843
20.9997 | -106.3874
21.0035 | -106.387
21.0073 | -106.3845
21.0111 | -106.3848
21.0148 | -106.3871
21.0186 | -106.3899
21.0224 | -106.3911
21.0262 | -106.3923
21.03 | -106.3918
21.0337 | -106.3916
21.0375 | -106.3921
21.0413 | -106.391
21.0451 | -106.3917
21.0489 | -106.3924
21.0526 | -106.3946
21.0564 | -106.3969
21.0602 | -106.3988
21.064 | -106.4011
21.0678 | -106.4041
21.0715 | -106.41
21.0753 | -106.4167
21.0791 | -106.4217
21.0829 | -106.4266
21.0866 | -106.4297
21.0904 | -106.4321
21.0942 | -106.4357
21.098 | -106.4416
21.1018 | -106.4452
21.1055 | -106.4483
21.1093 | -106.4524
21.1131 | -106.4561
21.1169 | -106.4587
21.1207 | -106.4623
21.1244 | -106.4649
21.1282 | -106.4657
21.132 | -106.466
21.1358 | -106.4671
21.1396 | -106.4697
21.1433 | -106.4701
21.1471 | -106.4691
21.1509 | -106.4669
21.1547 | -106.4637
21.1585 | -106.459
21.1622 | -106.4555
21.166 | -106.4523
21.1698 | -106.4484
21.1736 | -106.4457
21.1774 | -106.4414
21.1811 | -106.4345
21.1849 | -106.4266
21.1887 | -106.4204
21.1925 | -106.4145
21.1963 | -106.4058
21.2 | -106.3959
21.2038 | -106.3875
21.2076 | -106.3806
21.2114 | -106.3714
21.2152 | -106.361
21.2189 | -106.3504
21.2227 | -106.3397
21.2265 | -106.3311
21.2303 | -106.3241
21.2341 | -106.318
21.2378 | -106.3119
21.2416 | -106.3075
21.2454 | -106.3003
21.2492 | -106.2937
21.253 | -106.2907
21.2567 | -106.2868
21.2605 | -106.2833
21.2643 | -106.2876
21.2681 | -106.2859
21.2719 | -106.2852
21.2756 | -106.2868
21.2794 | -106.2891
21.2832 | -106.2918
21.287 | -106.2958
21.2907 | -106.3007
21.2945 | -106.3079
21.2983 | -106.3166
21.3021 | -106.3227
21.3059 | -106.3302
21.3096 | -106.3399
21.3134 | -106.3481
21.3172 | -106.3587
21.321 | -106.3685
21.3248 | -106.3797
21.3285 | -106.3928
21.3323 | -106.4059
21.3361 | -106.4179
21.3399 | -106.4317
21.3437 | -106.4467
21.3474 | -106.4617
21.3512 | -106.4738
21.355 | -106.4869
21.3588 | -106.5018
21.3626 | -106.5169
21.3663 | -106.5321
21.3701 | -106.5481
21.3739 | -106.5613
21.3777 | -106.5757
21.3815 | -106.5918
21.3852 | -106.6056
21.389 | -106.6191
21.3928 | -106.6332
21.3966 | -106.648
21.4004 | -106.6607
21.4041 | -106.6726
21.4079 | -106.6845
21.4117 | -106.6945
21.4155 | -106.7018
21.4193 | -106.7095
21.423 | -106.7182
21.4268 | -106.7255
21.4306 | -106.7313
21.4344 | -106.736
21.4382 | -106.7394
21.4419 | -106.742
21.4457 | -106.7416
21.4495 | -106.7408
21.4533 | -106.741
21.4571 | -106.7376
21.4608 | -106.7333
21.4646 | -106.7266
21.4684 | -106.719
21.4722 | -106.7107
21.476 | -106.7007
21.4797 | -106.6883
21.4835 | -106.6771
21.4873 | -106.6643
21.4911 | -106.6521
21.4948 | -106.6384
21.4986 | -106.6249
21.5024 | -106.6104
21.5062 | -106.5936
21.51 | -106.5762
21.5137 | -106.5616
21.5175 | -106.5487
21.5213 | -106.5354
21.5251 | -106.5226
21.5289 | -106.508
21.5326 | -106.4912
21.5364 | -106.4772
21.5402 | -106.4653
21.544 | -106.4529
21.5478 | -106.4437
21.5515 | -106.4373
21.5553 | -106.4288
21.5591 | -106.4206
21.5629 | -106.414
21.5667 | -106.4073
21.5704 | -106.4006
21.5742 | -106.3949
21.578 | -106.3889
21.5818 | -106.3842
21.5856 | -106.3813
21.5893 | -106.3781
21.5931 | -106.3745
21.5969 | -106.3719
21.6007 | -106.369
21.6045 | -106.368
21.6082 | -106.3689
21.612 | -106.3689
21.6158 | -106.366
21.6196 | -106.366
21.6234 | -106.3646
21.6271 | -106.3634
21.6309 | -106.3619
21.6347 | -106.3588
21.6385 | -106.3552
21.6423 | -106.353
21.646 | -106.3492
21.6498 | -106.3443
21.6536 | -106.338
21.6574 | -106.3332
21.6612 | -106.3269
21.6649 | -106.3191
21.6687 | -106.3105
21.6725 | -106.3023
21.6763 | -106.2921
21.6801 | -106.282
21.6838 | -106.2713
21.6876 | -106.259
21.6914 | -106.2495
21.6952 | -106.2407
21.699 | -106.2292
21.7027 | -106.2167
21.7065 | -106.2053
21.7103 | -106.1949
21.7141 | -106.1841
21.7178 | -106.1727
21.7216 | -106.1607
21.7254 | -106.1487
21.7292 | -106.1365
21.733 | -106.1253
21.7367 | -106.1156
21.7405 | -106.1081
21.7443 | -106.1018
21.7481 | -106.0923
21.7519 | -106.0845
21.7556 | -106.0775
21.7594 | -106.0712
21.7632 | -106.0675
21.767 | -106.0647
21.7708 | -106.0638
21.7745 | -106.0626
21.7783 | -106.062
21.7821 | -106.0632
21.7859 | -106.0678
21.7897 | -106.0709
21.7934 | -106.0744
21.7972 | -106.0783
21.801 | -106.0844
21.8048 | -106.0915
21.8086 | -106.1057
21.8123 | -106.116
21.8161 | -106.1258
21.8199 | -106.1348
21.8237 | -106.1441
21.8275 | -106.1529
21.8312 | -106.1631
21.835 | -106.1729
21.8388 | -106.1838
21.8426 | -106.1962
21.8464 | -106.2076
21.8501 | -106.2177
21.8539 | -106.2282
21.8577 | -106.2402
21.8615 | -106.2499
21.8653 | -106.2588
21.869 | -106.2677
21.8728 | -106.2766
21.8766 | -106.2856
21.8804 | -106.2956
21.8842 | -106.3035
21.8879 | -106.3109
21.8917 | -106.317
21.8955 | -106.3218
21.8993 | -106.3266
21.9031 | -106.3308
21.9068 | -106.3347
21.9106 | -106.3379
21.9144 | -106.3387
21.9182 | -106.3393
21.9219 | -106.3392
21.9257 | -106.3383
21.9295 | -106.3314
21.9333 | -106.3236
21.9371 | -106.3171
21.9408 | -106.3105
21.9446 | -106.3015
21.9484 | -106.2936
21.9522 | -106.2843
21.956 | -106.2748
21.9597 | -106.2645
21.9635 | -106.2503
21.9673 | -106.2381
21.9711 | -106.2259
21.9749 | -106.2162
21.9786 | -106.2057
21.9824 | -106.1922
21.9862 | -106.1792
21.99 | -106.1673
21.9938 | -106.1557
21.9975 | -106.1444
22.0013 | -106.1338
22.0051 | -106.1233
22.0089 | -106.1134
22.0127 | -106.1058
22.0164 | -106.098
22.0202 | -106.0907
22.024 | -106.0819
22.0278 | -106.0756
22.0316 | -106.0716
22.0353 | -106.0675
22.0391 | -106.0643
22.0429 | -106.0631
22.0467 | -106.0614
22.0505 | -106.0582
22.0542 | -106.0562
22.058 | -106.0578
22.0618 | -106.0597
22.0656 | -106.0616
22.0694 | -106.0639
22.0731 | -106.067
22.0769 | -106.0699
22.0807 | -106.0717
22.0845 | -106.0762
22.0883 | -106.0813
22.092 | -106.085
22.0958 | -106.0871
22.0996 | -106.0922
22.1034 | -106.0978
22.1072 | -106.1017
22.1109 | -106.1049
22.1147 | -106.1099
22.1185 | -106.112
22.1223 | -106.1142
22.1261 | -106.1181
22.1298 | -106.1218
22.1336 | -106.1259
22.1374 | -106.13
22.1412 | -106.1351
22.1449 | -106.139
22.1487 | -106.1447
22.1525 | -106.1494
22.1563 | -106.1525
22.1601 | -106.1561
22.1638 | -106.1611
22.1676 | -106.1649
22.1714 | -106.1682
22.1752 | -106.1737
22.179 | -106.178
22.1827 | -106.1818
22.1865 | -106.1848
22.1903 | -106.1877
22.1941 | -106.1894
22.1979 | -106.193
22.2016 | -106.1961
22.2054 | -106.1995
22.2092 | -106.2041
22.213 | -106.2079
22.2168 | -106.2104
22.2205 | -106.2131
22.2243 | -106.216
22.2281 | -106.2192
22.2319 | -106.2235
22.2357 | -106.2271
22.2394 | -106.2326
22.2432 | -106.2393
22.247 | -106.2449
22.2508 | -106.2495
22.2546 | -106.2536
22.2583 | -106.259
22.2621 | -106.265
22.2659 | -106.2697
22.2697 | -106.2731
22.2735 | -106.2761
22.2772 | -106.2828
22.281 | -106.2853
22.2848 | -106.2882
22.2886 | -106.2899
22.2924 | -106.292
22.2961 | -106.2954
22.2999 | -106.299
22.3037 | -106.302
22.3075 | -106.3031
22.3113 | -106.3053
22.315 | -106.3047
22.3188 | -106.3056
22.3226 | -106.3072
22.3264 | -106.3069
22.3302 | -106.3056
22.3339 | -106.3048
22.3377 | -106.3047
22.3415 | -106.3032
22.3453 | -106.3008
22.349 | -106.297
22.3528 | -106.2922
22.3566 | -106.2828
22.3604 | -106.2738
22.3642 | -106.267
22.3679 | -106.2624
22.3717 | -106.2564
22.3755 | -106.2488
22.3793 | -106.2398
22.3831 | -106.2297
22.3868 | -106.2207
22.3906 | -106.2116
22.3944 | -106.2034
22.3982 | -106.1944
22.402 | -106.1833
22.4057 | -106.1725
22.4095 | -106.1616
22.4133 | -106.1518
22.4171 | -106.1406
22.4209 | -106.13
22.4246 | -106.1201
22.4284 | -106.1108
22.4322 | -106.1005
22.436 | -106.0932
22.4398 | -106.0854
22.4435 | -106.0749
22.4473 | -106.0659
22.4511 | -106.0592
22.4549 | -106.0512
22.4587 | -106.0412
22.4624 | -106.0324
22.4662 | -106.0225
22.47 | -106.0135
22.4738 | -106.0064
22.4776 | -105.9997
22.4813 | -105.9909
22.4851 | -105.9837
22.4889 | -105.9744
22.4927 | -105.9644
22.4965 | -105.9589
22.5002 | -105.9544
22.504 | -105.9485
22.5078 | -105.9424
22.5116 | -105.9368
22.5154 | -105.9303
22.5191 | -105.9256
22.5229 | -105.9195
22.5267 | -105.9139
22.5305 | -105.9102
22.5343 | -105.9064
22.538 | -105.9021
22.5418 | -105.8969
22.5456 | -105.8901
22.5494 | -105.8826
22.5532 | -105.8774
22.5569 | -105.8713
22.5607 | -105.8642
22.5645 | -105.8589
22.5683 | -105.8527
22.5721 | -105.8457
22.5758 | -105.8379
22.5796 | -105.8296
22.5834 | -105.8224
22.5872 | -105.8155
22.5909 | -105.8086
22.5947 | -105.7999
22.5985 | -105.7927
22.6023 | -105.7835
22.6061 | -105.7749
22.6098 | -105.7666
22.6136 | -105.7558
22.6174 | -105.7449
22.6212 | -105.7342
22.625 | -105.7246
22.6287 | -105.7135
22.6325 | -105.7031
22.6363 | -105.6923
22.6401 | -105.6812
22.6439 | -105.6685
22.6476 | -105.6565
22.6514 | -105.6447
22.6552 | -105.6336
22.659 | -105.6236
22.6628 | -105.613
22.6665 | -105.6008
22.6703 | -105.589
22.6741 | -105.5783
22.6779 | -105.5694
22.6817 | -105.5614
22.6854 | -105.5528
22.6892 | -105.5412
22.693 | -105.5322
22.6968 | -105.5238
22.7006 | -105.5175
22.7043 | -105.5116
22.7081 | -105.5045
22.7119 | -105.4981
22.7157 | -105.4932
22.7195 | -105.4903
22.7232 | -105.4883
22.727 | -105.4869
22.7308 | -105.4849
22.7346 | -105.483
22.7384 | -105.4837
22.7421 | -105.4845
22.7459 | -105.4851
22.7497 | -105.4862
22.7535 | -105.4879
22.7573 | -105.4892
22.761 | -105.4919
22.7648 | -105.4966
22.7686 | -105.5019
22.7724 | -105.5065
22.7762 | -105.5129
22.7799 | -105.5195
22.7837 | -105.5272
22.7875 | -105.5379
22.7913 | -105.5491
22.795 | -105.5591
22.7988 | -105.5689
22.8026 | -105.5798
22.8064 | -105.5911
22.8102 | -105.604
22.8139 | -105.6182
22.8177 | -105.6296
22.8215 | -105.6394
22.8253 | -105.6511
22.8291 | -105.6625
22.8328 | -105.6746
22.8366 | -105.6882
22.8404 | -105.7025
22.8442 | -105.7166
22.848 | -105.7288
22.8517 | -105.741
22.8555 | -105.7534
22.8593 | -105.7657
22.8631 | -105.781
22.8669 | -105.7984
22.8706 | -105.8142
22.8744 | -105.8282
22.8782 | -105.8416
22.882 | -105.8527
22.8858 | -105.8641
22.8895 | -105.873
22.8933 | -105.8824
22.8971 | -105.8938
22.9009 | -105.9008
22.9047 | -105.9097
22.9084 | -105.9177
22.9122 | -105.9245
22.916 | -105.9304
22.9198 | -105.938
22.9236 | -105.9444
22.9273 | -105.9503
22.9311 | -105.9534
22.9349 | -105.9549
22.9387 | -105.9578
22.9425 | -105.9592
22.9462 | -105.9589
22.95 | -105.9563
22.9538 | -105.9525
22.9576 | -105.9491
22.9614 | -105.9433
22.9651 | -105.9366
22.9689 | -105.9301
22.9727 | -105.9222
22.9765 | -105.9141
22.9803 | -105.9037
22.984 | -105.8893
22.9878 | -105.877
22.9916 | -105.8668
22.9954 | -105.854
22.9992 | -105.8408
23.0029 | -105.8261
23.0067 | -105.808
23.0105 | -105.7909
23.0143 | -105.7747
23.018 | -105.7567
23.0218 | -105.7403
23.0256 | -105.7219
23.0294 | -105.7025
23.0332 | -105.6852
23.0369 | -105.6672
23.0407 | -105.6485
23.0445 | -105.6305
23.0483 | -105.6134
23.0521 | -105.5958
23.0558 | -105.5812
23.0596 | -105.5658
23.0634 | -105.5481
23.0672 | -105.5328
23.071 | -105.5177
23.0747 | -105.5015
23.0785 | -105.4861
23.0823 | -105.472
23.0861 | -105.4576
23.0899 | -105.4434
23.0936 | -105.4303
23.0974 | -105.4172
23.1012 | -105.4052
23.105 | -105.3949
23.1088 | -105.3844
23.1125 | -105.372
23.1163 | -105.3598
23.1201 | -105.351
23.1239 | -105.3436
23.1277 | -105.3359
23.1314 | -105.3288
23.1352 | -105.3236
23.139 | -105.3179
23.1428 | -105.3128
23.1466 | -105.3096
23.1503 | -105.3075
23.1541 | -105.3058
23.1579 | -105.3017
23.1617 | -105.2983
23.1655 | -105.2963
23.1692 | -105.2937
23.173 | -105.2931
23.1768 | -105.2925
23.1806 | -105.2925
23.1844 | -105.2936
23.1881 | -105.2934
23.1919 | -105.293
23.1957 | -105.2941
23.1995 | -105.2949
23.2033 | -105.2952
23.207 | -105.2973
23.2108 | -105.2989
23.2146 | -105.3023
23.2184 | -105.3049
23.2221 | -105.3047
23.2259 | -105.3058
23.2297 | -105.3087
23.2335 | -105.3116
23.2373 | -105.3147
23.241 | -105.3196
23.2448 | -105.3247
23.2486 | -105.3284
23.2524 | -105.3317
23.2562 | -105.3362
23.2599 | -105.3413
23.2637 | -105.3469
23.2675 | -105.3527
23.2713 | -105.3597
23.2751 | -105.3669
23.2788 | -105.3743
23.2826 | -105.3823
23.2864 | -105.3902
23.2902 | -105.3985
23.294 | -105.4078
23.2977 | -105.4167
23.3015 | -105.4263
23.3053 | -105.4353
23.3091 | -105.4453
23.3129 | -105.4562
23.3166 | -105.4635
23.3204 | -105.4714
23.3242 | -105.4789
23.328 | -105.4859
23.3318 | -105.4946
23.3355 | -105.5055
23.3393 | -105.5149
23.3431 | -105.5247
23.3469 | -105.5324
23.3507 | -105.5374
23.3544 | -105.5447
23.3582 | -105.5533
23.362 | -105.56
23.3658 | -105.5669
23.3696 | -105.5747
23.3733 | -105.5818
23.3771 | -105.5886
23.3809 | -105.5949
23.3847 | -105.5993
23.3885 | -105.6067
23.3922 | -105.6152
23.396 | -105.6216
23.3998 | -105.6284
23.4036 | -105.636
23.4074 | -105.6435
23.4111 | -105.6508
23.4149 | -105.6571
23.4187 | -105.6633
23.4225 | -105.6691
23.4262 | -105.6763
23.43 | -105.6815
23.4338 | -105.6871
23.4376 | -105.6932
23.4414 | -105.6987
23.4451 | -105.7021
23.4489 | -105.7077
23.4527 | -105.7108
23.4565 | -105.7114
23.4603 | -105.7129
23.464 | -105.7141
23.4678 | -105.7138
23.4716 | -105.714
23.4754 | -105.7151
23.4792 | -105.7148
23.4829 | -105.7128
23.4867 | -105.7126
23.4905 | -105.714
23.4943 | -105.7146
23.4981 | -105.7136
23.5018 | -105.7135
23.5056 | -105.7108
23.5094 | -105.7091
23.5132 | -105.7091
23.517 | -105.7079
23.5207 | -105.7048
23.5245 | -105.7027
23.5283 | -105.7003
23.5321 | -105.6986
23.5359 | -105.6976
23.5396 | -105.6954
23.5434 | -105.6921
23.5472 | -105.6877
23.551 | -105.6847
23.5548 | -105.683
23.5585 | -105.6797
23.5623 | -105.6749
23.5661 | -105.6707
23.5699 | -105.6676
23.5737 | -105.6642
23.5774 | -105.6611
23.5812 | -105.6594
23.585 | -105.6576
23.5888 | -105.6554
23.5926 | -105.6538
23.5963 | -105.653
23.6001 | -105.652
23.6039 | -105.6504
23.6077 | -105.6493
23.6115 | -105.65
23.6152 | -105.6509
23.619 | -105.6506
23.6228 | -105.6494
23.6266 | -105.6495
23.6304 | -105.6524
23.6341 | -105.6539
23.6379 | -105.6556
23.6417 | -105.6581
23.6455 | -105.6624
23.6492 | -105.668
23.653 | -105.6719
23.6568 | -105.6778
23.6606 | -105.6842
23.6644 | -105.69
23.6681 | -105.6964
23.6719 | -105.7026
23.6757 | -105.7097
23.6795 | -105.7181
23.6833 | -105.727
23.687 | -105.7351
23.6908 | -105.7425
23.6946 | -105.7513
23.6984 | -105.7609
23.7022 | -105.7687
23.7059 | -105.7767
23.7097 | -105.7854
23.7135 | -105.7951
23.7173 | -105.8064
23.7211 | -105.8139
23.7248 | -105.821
23.7286 | -105.8279
23.7324 | -105.8358
23.7362 | -105.8437
23.74 | -105.8514
23.7437 | -105.8578
23.7475 | -105.8662
23.7513 | -105.8755
23.7551 | -105.8845
23.7589 | -105.8909
23.7626 | -105.8992
23.7664 | -105.9077
23.7702 | -105.9144
23.774 | -105.9232
23.7778 | -105.9327
23.7815 | -105.9413
23.7853 | -105.9507
23.7891 | -105.9613
23.7929 | -105.9724
23.7967 | -105.9831
23.8004 | -105.9941
23.8042 | -106.0048
23.808 | -106.0132
23.8118 | -106.0212
23.8156 | -106.0321
23.8193 | -106.0437
23.8231 | -106.0518
23.8269 | -106.0604
23.8307 | -106.0701
23.8345 | -106.0779
23.8382 | -106.0872
23.842 | -106.0984
23.8458 | -106.1089
23.8496 | -106.117
23.8533 | -106.1265
23.8571 | -106.1374
23.8609 | -106.1464
23.8647 | -106.1548
23.8685 | -106.1636
23.8722 | -106.1724
23.876 | -106.1813
23.8798 | -106.1884
23.8836 | -106.1957
23.8874 | -106.2032
23.8911 | -106.209
23.8949 | -106.2137
23.8987 | -106.2183
23.9025 | -106.2238
23.9063 | -106.2277
23.91 | -106.2305
23.9138 | -106.2331
23.9176 | -106.235
23.9214 | -106.237
23.9252 | -106.2384
23.9289 | -106.2397
23.9327 | -106.2405
23.9365 | -106.241
23.9403 | -106.2422
23.9441 | -106.2431
23.9478 | -106.2419
23.9516 | -106.2416
23.9554 | -106.2419
23.9592 | -106.2433
23.963 | -106.2435
23.9667 | -106.2422
23.9705 | -106.242
23.9743 | -106.2426
23.9781 | -106.2422
23.9819 | -106.2421
23.9856 | -106.2428
23.9894 | -106.2474
23.9932 | -106.2484
23.997 | -106.2512
24.0008 | -106.2541
24.0045 | -106.257
24.0083 | -106.2597
24.0121 | -106.2638
24.0159 | -106.2681
24.0197 | -106.2727
24.0234 | -106.2788
24.0272 | -106.2857
24.031 | -106.2925
24.0348 | -106.2992
24.0386 | -106.3057
24.0423 | -106.314
24.0461 | -106.3233
24.0499 | -106.3315
24.0537 | -106.3387
24.0575 | -106.3475
24.0612 | -106.3586
24.065 | -106.3687
24.0688 | -106.3795
24.0726 | -106.3887
24.0763 | -106.3984
24.0801 | -106.4092
24.0839 | -106.4221
24.0877 | -106.4337
24.0915 | -106.4444
24.0952 | -106.4557
24.099 | -106.4682
24.1028 | -106.48
24.1066 | -106.4897
24.1104 | -106.5007
24.1141 | -106.5105
24.1179 | -106.5208
24.1217 | -106.531
24.1255 | -106.5414
24.1293 | -106.5522
24.133 | -106.562
24.1368 | -106.572
24.1406 | -106.5801
24.1444 | -106.5893
24.1482 | -106.6004
24.1519 | -106.6118
24.1557 | -106.6193
24.1595 | -106.6274
24.1633 | -106.6384
24.1671 | -106.649
24.1708 | -106.6551
24.1746 | -106.6626
24.1784 | -106.6692
24.1822 | -106.6732
24.186 | -106.6765
24.1897 | -106.68
24.1935 | -106.683
24.1973 | -106.6865
24.2011 | -106.6881
24.2049 | -106.6895
24.2086 | -106.6893
24.2124 | -106.6882
24.2162 | -106.6876
24.22 | -106.686
24.2238 | -106.6855
24.2275 | -106.6834
24.2313 | -106.6816
24.2351 | -106.6797
24.2389 | -106.6773
24.2427 | -106.6727
24.2464 | -106.6672
24.2502 | -106.6628
24.254 | -106.6576
24.2578 | -106.6512
24.2616 | -106.646
24.2653 | -106.6412
24.2691 | -106.6346
24.2729 | -106.6262
24.2767 | -106.6151
24.2804 | -106.6028
24.2842 | -106.5844
24.288 | -106.5619
24.2918 | -106.5379
24.2956 | -106.5085
24.2993 | -106.4797
24.3031 | -106.4548
24.3069 | -106.4295
24.3107 | -106.4019
24.3145 | -106.3718
24.3182 | -106.3434
24.322 | -106.3147
24.3258 | -106.2837
24.3296 | -106.2516
24.3334 | -106.2179
24.3371 | -106.185
24.3409 | -106.1493
24.3447 | -106.1125
24.3485 | -106.077
24.3523 | -106.0405
24.356 | -106.0031
24.3598 | -105.9645
24.3636 | -105.9242
24.3674 | -105.882
24.3712 | -105.8408
24.3749 | -105.8009
24.3787 | -105.7593
24.3825 | -105.7192
24.3863 | -105.6792
24.3901 | -105.6384
24.3938 | -105.6
24.3976 | -105.5624
24.4014 | -105.5248
24.4052 | -105.4886
24.409 | -105.4532
24.4127 | -105.4193
24.4165 | -105.3873
24.4203 | -105.3569
24.4241 | -105.328
24.4279 | -105.3011
24.4316 | -105.2764
24.4354 | -105.2545
24.4392 | -105.2339
24.443 | -105.2167
24.4468 | -105.2021
24.4505 | -105.1919
24.4543 | -105.1842
24.4581 | -105.1783
24.4619 | -105.1756
24.4657 | -105.1748
24.4694 | -105.1761
24.4732 | -105.1796
24.477 | -105.188
24.4808 | -105.1987
24.4846 | -105.2123
24.4883 | -105.2262
24.4921 | -105.2411
24.4959 | -105.2587
24.4997 | -105.2765
24.5034 | -105.2967
24.5072 | -105.3193
24.511 | -105.3436
24.5148 | -105.3713
24.5186 | -105.4003
24.5223 | -105.4302
24.5261 | -105.4612
24.5299 | -105.4927
24.5337 | -105.5242
24.5375 | -105.5549
24.5412 | -105.5844
24.545 | -105.6142
24.5488 | -105.6458
24.5526 | -105.6769
24.5564 | -105.7066
24.5601 | -105.7334
24.5639 | -105.7586
24.5677 | -105.7814
24.5715 | -105.792
24.5753 | -105.7879
24.579 | -105.7671
24.5828 | -105.7178
24.5866 | -105.6681
24.5904 | -105.6177
24.5942 | -105.5608
24.5979 | -105.5022
24.6017 | -105.4411
24.6055 | -105.3824
24.6093 | -105.3207
24.6131 | -105.2561
24.6168 | -105.1919
24.6206 | -105.1244
24.6244 | -105.0529
24.6282 | -104.9743
24.632 | -104.8944
24.6357 | -104.8134
24.6395 | -104.7269
24.6433 | -104.639
24.6471 | -104.5454
24.6509 | -104.4443
24.6546 | -104.3401
24.6584 | -104.2358
24.6622 | -104.1309
24.666 | -104.0229
24.6698 | -103.9143
24.6735 | -103.8063
24.6773 | -103.6982
24.6811 | -103.5937
24.6849 | -103.4903
24.6887 | -103.3898
24.6924 | -103.2917
24.6962 | -103.1958
24.7 | -103.1037
24.7038 | -103.0149
24.7076 | -102.9303
24.7113 | -102.8507
24.7151 | -102.7763
24.7189 | -102.7062
24.7227 | -102.6408
24.7264 | -102.5815
24.7302 | -102.5279
24.734 | -102.4803
24.7378 | -102.4388
24.7416 | -102.4025
24.7453 | -102.3726
24.7491 | -102.3491
24.7529 | -102.3319
24.7567 | -102.3214
24.7605 | -102.318
24.7642 | -102.3209
24.768 | -102.3309
24.7718 | -102.347
24.7756 | -102.3688
24.7794 | -102.3975
24.7831 | -102.4327
24.7869 | -102.4749
24.7907 | -102.5229
24.7945 | -102.5769
24.7983 | -102.6375
24.802 | -102.7032
24.8058 | -102.7748
24.8096 | -102.8518
24.8134 | -102.9337
24.8172 | -103.0208
24.8209 | -103.1123
24.8247 | -103.2073
24.8285 | -103.3065
24.8323 | -103.4096
24.8361 | -103.5161
24.8398 | -103.6264
24.8436 | -103.7374
24.8474 | -103.8516
24.8512 | -103.9665
24.855 | -104.0816
24.8587 | -104.1968
24.8625 | -104.3108
24.8663 | -104.4244
24.8701 | -104.5385
24.8739 | -104.6508
24.8776 | -104.7605
24.8814 | -104.869
24.8852 | -104.973
24.889 | -105.0737
24.8928 | -105.1717
24.8965 | -105.2661
24.9003 | -105.3552
24.9041 | -105.4388
24.9079 | -105.5192
24.9117 | -105.5975
24.9154 | -105.6713
24.9192 | -105.7412
24.923 | -105.813
24.9268 | -105.8899
24.9305 | -105.9734
24.9343 | -106.0714
24.9381 | -106.1507
24.9419 | -106.215
24.9457 | -106.27
24.9494 | -106.3095
24.9532 | -106.3391
24.957 | -106.3554
24.9608 | -106.3642
24.9646 | -106.3661
24.9683 | -106.3614
24.9721 | -106.354
24.9759 | -106.3478
24.9797 | -106.3431
24.9835 | -106.3375
24.9872 | -106.3307
24.991 | -106.3224
24.9948 | -106.3113
24.9986 | -106.3019
25.0024 | -106.2942
25.0061 | -106.2871
25.0099 | -106.2785
25.0137 | -106.2715
25.0175 | -106.2648
25.0213 | -106.2554
25.025 | -106.2467
25.0288 | -106.2412
25.0326 | -106.2336
25.0364 | -106.226
25.0402 | -106.2175
25.0439 | -106.2101
25.0477 | -106.2033
25.0515 | -106.1976
25.0553 | -106.1938
25.0591 | -106.1901
25.0628 | -106.1865
25.0666 | -106.1833
25.0704 | -106.1795
25.0742 | -106.176
25.078 | -106.1734
25.0817 | -106.1751
25.0855 | -106.1728
25.0893 | -106.1696
25.0931 | -106.169
25.0969 | -106.1697
25.1006 | -106.1707
25.1044 | -106.1704
25.1082 | -106.1709
25.112 | -106.1725
25.1158 | -106.1724
25.1195 | -106.1735
25.1233 | -106.1746
25.1271 | -106.1776
25.1309 | -106.181
25.1347 | -106.1834
25.1384 | -106.1858
25.1422 | -106.1877
25.146 | -106.1905
25.1498 | -106.1941
25.1535 | -106.195
25.1573 | -106.1984
25.1611 | -106.2034
25.1649 | -106.2078
25.1687 | -106.2136
25.1724 | -106.2175
25.1762 | -106.222
25.18 | -106.2257
25.1838 | -106.2289
25.1876 | -106.2322
25.1913 | -106.2363
25.1951 | -106.2392
25.1989 | -106.2429
25.2027 | -106.2457
25.2065 | -106.2479
25.2102 | -106.2499
25.214 | -106.2515
25.2178 | -106.2533
25.2216 | -106.2554
25.2254 | -106.2571
25.2291 | -106.2593
25.2329 | -106.2596
25.2367 | -106.2588
25.2405 | -106.258
25.2443 | -106.2565
25.248 | -106.2557
25.2518 | -106.2534
25.2556 | -106.2507
25.2594 | -106.2469
25.2632 | -106.2422
25.2669 | -106.2355
25.2707 | -106.2271
25.2745 | -106.2193
25.2783 | -106.2114
25.2821 | -106.2022
25.2858 | -106.1906
25.2896 | -106.1782
25.2934 | -106.1666
25.2972 | -106.1566
25.301 | -106.1451
25.3047 | -106.134
25.3085 | -106.1222
25.3123 | -106.1095
25.3161 | -106.0969
25.3199 | -106.0853
25.3236 | -106.0729
25.3274 | -106.0608
25.3312 | -106.0464
25.335 | -106.0333
25.3388 | -106.0209
25.3425 | -106.0069
25.3463 | -105.996
25.3501 | -105.9861
25.3539 | -105.9734
25.3576 | -105.9598
25.3614 | -105.9468
25.3652 | -105.9354
25.369 | -105.9241
25.3728 | -105.9124
25.3765 | -105.9038
25.3803 | -105.8958
25.3841 | -105.8864
25.3879 | -105.8773
25.3917 | -105.8689
25.3954 | -105.8616
25.3992 | -105.8555
25.403 | -105.8492
25.4068 | -105.842
25.4106 | -105.8346
25.4143 | -105.8301
25.4181 | -105.8276
25.4219 | -105.8249
25.4257 | -105.8233
25.4295 | -105.8225
25.4332 | -105.8206
25.437 | -105.82
25.4408 | -105.8187
25.4446 | -105.8194
25.4484 | -105.8222
25.4521 | -105.8254
25.4559 | -105.8282
25.4597 | -105.8318
25.4635 | -105.8365
25.4673 | -105.8413
25.471 | -105.845
25.4748 | -105.8487
25.4786 | -105.8552
25.4824 | -105.8619
25.4862 | -105.8689
25.4899 | -105.8748
25.4937 | -105.8815
25.4975 | -105.8893
25.5013 | -105.8975
25.5051 | -105.9061
25.5088 | -105.9135
25.5126 | -105.9226
25.5164 | -105.9305
25.5202 | -105.9374
25.524 | -105.9451
25.5277 | -105.9524
25.5315 | -105.9608
25.5353 | -105.9671
25.5391 | -105.9752
25.5429 | -105.9837
25.5466 | -105.9915
25.5504 | -105.9989
25.5542 | -106.0047
25.558 | -106.0098
25.5617 | -106.0161
25.5655 | -106.0221
25.5693 | -106.0273
25.5731 | -106.0333
25.5769 | -106.0389
25.5806 | -106.0443
25.5844 | -106.048
25.5882 | -106.0522
25.592 | -106.0571
25.5958 | -106.0625
25.5995 | -106.0675
25.6033 | -106.0718
25.6071 | -106.0768
25.6109 | -106.0821
25.6147 | -106.0871
25.6184 | -106.0932
25.6222 | -106.097
25.626 | -106.0997
25.6298 | -106.1027
25.6336 | -106.1071
25.6373 | -106.1104
25.6411 | -106.1139
25.6449 | -106.1164
25.6487 | -106.1196
25.6525 | -106.1229
25.6562 | -106.1249
25.66 | -106.1263
25.6638 | -106.1279
25.6676 | -106.1289
25.6714 | -106.1302
25.6751 | -106.1307
25.6789 | -106.132
25.6827 | -106.1341
25.6865 | -106.135
25.6903 | -106.1367
25.694 | -106.1394
25.6978 | -106.1415
25.7016 | -106.1438
25.7054 | -106.1471
25.7092 | -106.1473
25.7129 | -106.1485
25.7167 | -106.1522
25.7205 | -106.1553
25.7243 | -106.1567
25.7281 | -106.1581
25.7318 | -106.1573
25.7356 | -106.159
25.7394 | -106.1602
25.7432 | -106.1618
25.747 | -106.1636
25.7507 | -106.1644
25.7545 | -106.1656
25.7583 | -106.167
25.7621 | -106.1695
25.7659 | -106.1723
25.7696 | -106.1766
25.7734 | -106.1816
25.7772 | -106.1841
25.781 | -106.1856
25.7847 | -106.1867
25.7885 | -106.187
25.7923 | -106.1885
25.7961 | -106.1927
25.7999 | -106.1964
25.8036 | -106.2003
25.8074 | -106.2043
25.8112 | -106.2062
25.815 | -106.2083
25.8188 | -106.2127
25.8225 | -106.2167
25.8263 | -106.2206
25.8301 | -106.2242
25.8339 | -106.2284
25.8377 | -106.233
25.8414 | -106.2354
25.8452 | -106.2369
25.849 | -106.2394
25.8528 | -106.2435
25.8566 | -106.2473
25.8603 | -106.2496
25.8641 | -106.2527
25.8679 | -106.2567
25.8717 | -106.2592
25.8755 | -106.2598
25.8792 | -106.2599
25.883 | -106.2601
25.8868 | -106.2614
25.8906 | -106.263
25.8944 | -106.265
25.8981 | -106.2683
25.9019 | -106.2708
25.9057 | -106.2723
25.9095 | -106.2723
25.9133 | -106.2727
25.917 | -106.2756
25.9208 | -106.2774
25.9246 | -106.2784
25.9284 | -106.28
25.9322 | -106.28
25.9359 | -106.2804
25.9397 | -106.2814
25.9435 | -106.2802
25.9473 | -106.2803
25.9511 | -106.2803
25.9548 | -106.2789
25.9586 | -106.2771
25.9624 | -106.2766
25.9662 | -106.2779
25.97 | -106.278
25.9737 | -106.278
25.9775 | -106.2778
25.9813 | -106.278
25.9851 | -106.2786
25.9888 | -106.278
25.9926 | -106.2788
25.9964 | -106.2809
26.0002 | -106.2829
26.004 | -106.2852
26.0077 | -106.2869
26.0115 | -106.2894
26.0153 | -106.2914
26.0191 | -106.2923
26.0229 | -106.2931
26.0266 | -106.2929
26.0304 | -106.2941
26.0342 | -106.2944
26.038 | -106.2951
26.0418 | -106.2954
26.0455 | -106.2952
26.0493 | -106.2942
26.0531 | -106.2932
26.0569 | -106.2937
26.0607 | -106.2956
26.0644 | -106.297
26.0682 | -106.2981
26.072 | -106.2999
26.0758 | -106.3039
26.0796 | -106.3067
26.0833 | -106.3077
26.0871 | -106.3098
26.0909 | -106.3127
26.0947 | -106.315
26.0985 | -106.3198
26.1022 | -106.3236
26.106 | -106.3254
26.1098 | -106.3277
26.1136 | -106.3302
26.1174 | -106.3322
26.1211 | -106.3349
26.1249 | -106.3373
26.1287 | -106.3367
26.1325 | -106.3376
26.1363 | -106.3393
26.14 | -106.3407
26.1438 | -106.3424
26.1476 | -106.3437
26.1514 | -106.3459
26.1552 | -106.3476
26.1589 | -106.3476
26.1627 | -106.3469
26.1665 | -106.3462
26.1703 | -106.3414
26.1741 | -106.3415
26.1778 | -106.3396
26.1816 | -106.3374
26.1854 | -106.3356
26.1892 | -106.3329
26.193 | -106.3302
26.1967 | -106.3273
26.2005 | -106.3245
26.2043 | -106.3219
26.2081 | -106.3184
26.2118 | -106.3141
26.2156 | -106.3104
26.2194 | -106.3056
26.2232 | -106.3008
26.227 | -106.2967
26.2307 | -106.2929
26.2345 | -106.287
26.2383 | -106.2805
26.2421 | -106.2764
26.2459 | -106.2735
26.2496 | -106.2712
26.2534 | -106.2677
26.2572 | -106.2627
26.261 | -106.2568
26.2648 | -106.2505
26.2685 | -106.2452
26.2723 | -106.2405
26.2761 | -106.2375
26.2799 | -106.2353
26.2837 | -106.2323
26.2874 | -106.2297
26.2912 | -106.229
26.295 | -106.2261
26.2988 | -106.2232
26.3026 | -106.2216
26.3063 | -106.2197
26.3101 | -106.2182
26.3139 | -106.2184
26.3177 | -106.2171
26.3215 | -106.2163
26.3252 | -106.2196
26.329 | -106.2223
26.3328 | -106.2233
26.3366 | -106.2244
26.3404 | -106.2262
26.3441 | -106.2292
26.3479 | -106.2334
26.3517 | -106.2369
26.3555 | -106.2399
26.3593 | -106.2441
26.363 | -106.2492
26.3668 | -106.2546
26.3706 | -106.2585
26.3744 | -106.2633
26.3782 | -106.2702
26.3819 | -106.2776
26.3857 | -106.2854
26.3895 | -106.2934
26.3933 | -106.3033
26.3971 | -106.3138
26.4008 | -106.324
26.4046 | -106.3346
26.4084 | -106.3432
26.4122 | -106.3519
26.4159 | -106.3603
26.4197 | -106.3693
26.4235 | -106.3801
26.4273 | -106.3899
26.4311 | -106.3986
26.4348 | -106.4068
26.4386 | -106.4139
26.4424 | -106.421
26.4462 | -106.4264
26.45 | -106.4325
26.4537 | -106.4415
26.4575 | -106.4478
26.4613 | -106.4526
26.4651 | -106.458
26.4689 | -106.4625
26.4726 | -106.4679
26.4764 | -106.4718
26.4802 | -106.4735
26.484 | -106.4757
26.4878 | -106.4767
26.4915 | -106.4787
26.4953 | -106.4805
26.4991 | -106.4818
26.5029 | -106.4804
26.5067 | -106.4785
26.5104 | -106.4763
26.5142 | -106.474
26.518 | -106.4719
26.5218 | -106.4681
26.5256 | -106.4646
26.5293 | -106.4619
26.5331 | -106.4563
26.5369 | -106.4511
26.5407 | -106.4471
26.5445 | -106.4399
26.5482 | -106.4308
26.552 | -106.4236
26.5558 | -106.4177
26.5596 | -106.4091
26.5634 | -106.401
26.5671 | -106.3929
26.5709 | -106.3848
26.5747 | -106.3764
26.5785 | -106.3681
26.5823 | -106.3602
26.586 | -106.3534
26.5898 | -106.3466
26.5936 | -106.3392
26.5974 | -106.332
26.6012 | -106.324
26.6049 | -106.3152
26.6087 | -106.3077
26.6125 | -106.2993
26.6163 | -106.2899
26.6201 | -106.2812
26.6238 | -106.2726
26.6276 | -106.2649
26.6314 | -106.2593
26.6352 | -106.2544
26.6389 | -106.2491
26.6427 | -106.2435
26.6465 | -106.2375
26.6503 | -106.2324
26.6541 | -106.2271
26.6578 | -106.2222
26.6616 | -106.2174
26.6654 | -106.2125
26.6692 | -106.2106
26.673 | -106.2085
26.6767 | -106.2056
26.6805 | -106.2033
26.6843 | -106.2003
26.6881 | -106.1987
26.6919 | -106.1996
26.6956 | -106.2012
26.6994 | -106.2
26.7032 | -106.1986
26.707 | -106.199
26.7108 | -106.1992
26.7145 | -106.1977
26.7183 | -106.1996
26.7221 | -106.2007
26.7259 | -106.2027
26.7297 | -106.2051
26.7334 | -106.2088
26.7372 | -106.2122
26.741 | -106.2141
26.7448 | -106.2154
26.7486 | -106.2181
26.7523 | -106.2203
26.7561 | -106.2227
26.7599 | -106.225
26.7637 | -106.2277
26.7675 | -106.2318
26.7712 | -106.2365
26.775 | -106.2387
26.7788 | -106.2417
26.7826 | -106.2454
26.7864 | -106.248
26.7901 | -106.2519
26.7939 | -106.2563
26.7977 | -106.2588
26.8015 | -106.2609
26.8053 | -106.2634
26.809 | -106.2664
26.8128 | -106.2693
26.8166 | -106.2722
26.8204 | -106.2751
26.8242 | -106.279
26.8279 | -106.2829
26.8317 | -106.2854
26.8355 | -106.2895
26.8393 | -106.2946
26.8431 | -106.2973
26.8468 | -106.2987
26.8506 | -106.3005
26.8544 | -106.3035
26.8582 | -106.3058
26.8619 | -106.3081
26.8657 | -106.3119
26.8695 | -106.3119
26.8733 | -106.3124
26.8771 | -106.3136
26.8808 | -106.3157
26.8846 | -106.316
26.8884 | -106.3156
26.8922 | -106.3154
26.896 | -106.3144
26.8997 | -106.3148
26.9035 | -106.3151
26.9073 | -106.3141
26.9111 | -106.3142
26.9149 | -106.3126
26.9186 | -106.311
26.9224 | -106.3113
26.9262 | -106.3114
26.93 | -106.3099
26.9338 | -106.3072
26.9375 | -106.306
26.9413 | -106.3039
26.9451 | -106.3028
26.9489 | -106.3018
26.9527 | -106.301
26.9564 | -106.2997
26.9602 | -106.2974
26.964 | -106.2937
26.9678 | -106.2905
26.9716 | -106.2888
26.9753 | -106.2888
26.9791 | -106.2892
26.9829 | -106.2898
26.9867 | -106.2904
26.9905 | -106.2918
26.9942 | -106.2938
26.998 | -106.2961
27.0018 | -106.2967
27.0056 | -106.2965
27.0094 | -106.2951
27.0131 | -106.2968
27.0169 | -106.3003
27.0207 | -106.3013
27.0245 | -106.3029
27.0283 | -106.3052
27.032 | -106.3086
27.0358 | -106.3122
27.0396 | -106.3172
27.0434 | -106.3213
27.0472 | -106.325
27.0509 | -106.3293
27.0547 | -106.3354
27.0585 | -106.3416
27.0623 | -106.3469
27.066 | -106.3505
27.0698 | -106.353
27.0736 | -106.3561
27.0774 | -106.3611
27.0812 | -106.3672
27.0849 | -106.3732
27.0887 | -106.3778
27.0925 | -106.3853
27.0963 | -106.3921
27.1001 | -106.3966
27.1038 | -106.4007
27.1076 | -106.4048
27.1114 | -106.4102
27.1152 | -106.415
27.119 | -106.419
27.1227 | -106.4234
27.1265 | -106.4288
27.1303 | -106.4339
27.1341 | -106.4379
27.1379 | -106.4397
27.1416 | -106.4414
27.1454 | -106.443
27.1492 | -106.4442
27.153 | -106.4456
27.1568 | -106.4481
27.1605 | -106.4498
27.1643 | -106.4512
27.1681 | -106.4548
27.1719 | -106.4565
27.1757 | -106.4571
27.1794 | -106.4586
27.1832 | -106.4597
27.187 | -106.4595
27.1908 | -106.4588
27.1946 | -106.4586
27.1983 | -106.4598
27.2021 | -106.4595
27.2059 | -106.4594
27.2097 | -106.4573
27.2135 | -106.4539
27.2172 | -106.4518
27.221 | -106.4499
27.2248 | -106.4482
27.2286 | -106.4463
27.2324 | -106.4442
27.2361 | -106.4422
27.2399 | -106.44
27.2437 | -106.4361
27.2475 | -106.4323
27.2513 | -106.432
27.255 | -106.4304
27.2588 | -106.4272
27.2626 | -106.4232
27.2664 | -106.4206
27.2702 | -106.4169
27.2739 | -106.4128
27.2777 | -106.4091
27.2815 | -106.4065
27.2853 | -106.4037
27.289 | -106.402
27.2928 | -106.3987
27.2966 | -106.3945
27.3004 | -106.39
27.3042 | -106.3867
27.3079 | -106.3831
27.3117 | -106.3781
27.3155 | -106.3746
27.3193 | -106.3714
27.3231 | -106.3678
27.3268 | -106.3648
27.3306 | -106.3612
27.3344 | -106.3581
27.3382 | -106.3552
27.342 | -106.3536
27.3457 | -106.3533
27.3495 | -106.3531
27.3533 | -106.3514
27.3571 | -106.3504
27.3609 | -106.3478
27.3646 | -106.345
27.3684 | -106.3432
27.3722 | -106.3419
27.376 | -106.34
27.3798 | -106.3365
27.3835 | -106.3327
27.3873 | -106.3317
27.3911 | -106.3308
27.3949 | -106.3309
27.3987 | -106.3291
27.4024 | -106.3252
27.4062 | -106.3235
27.41 | -106.3223
27.4138 | -106.3212
27.4176 | -106.321
27.4213 | -106.3191
27.4251 | -106.3172
27.4289 | -106.3161
27.4327 | -106.315
27.4365 | -106.3138
27.4402 | -106.3114
27.444 | -106.3092
27.4478 | -106.3082
27.4516 | -106.3079
27.4554 | -106.307
27.4591 | -106.3061
27.4629 | -106.3038
27.4667 | -106.3002
27.4705 | -106.2972
27.4743 | -106.2946
27.478 | -106.2899
27.4818 | -106.2827
27.4856 | -106.2784
27.4894 | -106.2749
27.4931 | -106.2719
27.4969 | -106.2683
27.5007 | -106.2637
27.5045 | -106.2578
27.5083 | -106.2523
27.512 | -106.246
27.5158 | -106.24
27.5196 | -106.2333
27.5234 | -106.2287
27.5272 | -106.2241
27.5309 | -106.217
27.5347 | -106.2114
27.5385 | -106.204
27.5423 | -106.1957
27.5461 | -106.1879
27.5498 | -106.1794
27.5536 | -106.1701
27.5574 | -106.1619
27.5612 | -106.1533
27.565 | -106.145
27.5687 | -106.1363
27.5725 | -106.1269
27.5763 | -106.1158
27.5801 | -106.105
27.5839 | -106.0926
27.5876 | -106.081
27.5914 | -106.07
27.5952 | -106.0596
27.599 | -106.0482
27.6028 | -106.0352
27.6065 | -106.0237
27.6103 | -106.0129
27.6141 | -106.0031
27.6179 | -105.991
27.6217 | -105.9776
27.6254 | -105.9663
27.6292 | -105.9546
27.633 | -105.9377
27.6368 | -105.921
27.6406 | -105.9056
27.6443 | -105.8896
27.6481 | -105.8708
27.6519 | -105.8507
27.6557 | -105.8291
27.6595 | -105.8028
27.6632 | -105.7716
27.667 | -105.728
27.6708 | -105.6737
27.6746 | -105.6171
27.6784 | -105.5442
27.6821 | -105.4571
27.6859 | -105.3576
27.6897 | -105.2522
27.6935 | -105.1464
27.6973 | -105.0335
27.701 | -104.9192
27.7048 | -104.8024
27.7086 | -104.6838
27.7124 | -104.5567
27.7161 | -104.4217
27.7199 | -104.2772
27.7237 | -104.1233
27.7275 | -103.9653
27.7313 | -103.8055
27.735 | -103.644
27.7388 | -103.4815
27.7426 | -103.3179
27.7464 | -103.1505
27.7502 | -102.9823
27.7539 | -102.8079
27.7577 | -102.6341
27.7615 | -102.4613
27.7653 | -102.2887
27.7691 | -102.1174
27.7728 | -101.9436
27.7766 | -101.7683
27.7804 | -101.5926
27.7842 | -101.4184
27.788 | -101.2443
27.7917 | -101.073
27.7955 | -100.9044
27.7993 | -100.7362
27.8031 | -100.5681
27.8069 | -100.3743
27.8106 | -100.2173
27.8144 | -100.0642
27.8182 | -99.9136
27.822 | -99.7654
27.8258 | -99.6204
27.8295 | -99.4783
27.8333 | -99.3388
27.8371 | -99.2007
27.8409 | -99.064
27.8447 | -98.9337
27.8484 | -98.8091
27.8522 | -98.6872
27.856 | -98.5652
27.8598 | -98.4468
27.8636 | -98.3299
27.8673 | -98.2143
27.8711 | -98.0991
27.8749 | -97.9901
27.8787 | -97.8766
27.8825 | -97.7674
27.8862 | -97.6579
27.89 | -97.5496
27.8938 | -97.4446
27.8976 | -97.3396
27.9014 | -97.231
27.9051 | -97.1228
27.9089 | -97.0199
27.9127 | -96.9198
27.9165 | -96.8182
27.9202 | -96.7166
27.924 | -96.6171
27.9278 | -96.5173
27.9316 | -96.4206
27.9354 | -96.3252
27.9391 | -96.223
27.9429 | -96.1222
27.9467 | -96.0228
27.9505 | -95.9232
27.9543 | -95.8256
27.958 | -95.7258
27.9618 | -95.6294
27.9656 | -95.5328
27.9694 | -95.435
27.9732 | -95.3385
27.9769 | -95.2411
27.9807 | -95.1444
27.9845 | -95.0481
27.9883 | -94.951
27.9921 | -94.8549
27.9958 | -94.7596
27.9996 | -94.6648
28.0034 | -94.5707