## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 217 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 215
Hood | 02-92B/P | 154
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 997.3 | hPa |  | PRC PRO 01
Humidity | 41.2 | %/Hr
Internal temperature of calibration bench | 23.96 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1009 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.2 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.03 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.03 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3098 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 13/04/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 16/04/2018
Engineer | Anthony HUE | Approval |  | 23/04/2018
CEO | Robert GLEMOT | Approval |  | 23/04/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5104 | 57.5911
0.012209 | -36.0568 | 38.475
0.018313 | -34.9993 | 28.2293
0.024417 | -34.5344 | 21.9788
0.030521 | -34.4191 | 18.2646
0.036626 | -34.2764 | 15.2184
0.04273 | -34.1852 | 12.8449
0.048834 | -34.1388 | 11.3326
0.054938 | -34.07 | 10.1912
0.061043 | -34.0057 | 9.0887
0.067147 | -33.9952 | 8.2553
0.073251 | -33.9964 | 7.487
0.079355 | -34.0158 | 6.8233
0.08546 | -34.0309 | 6.2268
0.091564 | -34.0259 | 5.7869
0.097668 | -34.0056 | 5.5501
0.10377 | -33.9813 | 5.3141
0.10988 | -33.9628 | 5.0423
0.11598 | -33.9535 | 4.7493
0.12209 | -33.9559 | 4.4627
0.12819 | -33.967 | 4.1752
0.13429 | -33.9789 | 3.8817
0.1404 | -33.9823 | 3.5947
0.1465 | -33.9845 | 3.3299
0.15261 | -33.9817 | 3.1052
0.15871 | -33.9752 | 2.9423
0.16482 | -33.967 | 2.843
0.17092 | -33.9648 | 2.8109
0.17702 | -33.9628 | 2.7435
0.18313 | -33.9618 | 2.6693
0.18923 | -33.9613 | 2.5626
0.19534 | -33.9597 | 2.4363
0.20144 | -33.958 | 2.2933
0.20754 | -33.9549 | 2.1726
0.21365 | -33.9538 | 2.058
0.21975 | -33.9534 | 1.9601
0.22586 | -33.9548 | 1.8719
0.23196 | -33.9577 | 1.7985
0.23807 | -33.9622 | 1.7342
0.24417 | -33.9663 | 1.6737
0.25027 | -33.9704 | 1.6099
0.25638 | -33.9727 | 1.5468
0.26248 | -33.9742 | 1.4792
0.26859 | -33.9747 | 1.4137
0.27469 | -33.9746 | 1.3424
0.2808 | -33.9752 | 1.2687
0.2869 | -33.9756 | 1.1975
0.293 | -33.9762 | 1.1306
0.29911 | -33.9768 | 1.0709
0.30521 | -33.9775 | 1.0119
0.31132 | -33.9783 | 0.97307
0.31742 | -33.9795 | 0.93052
0.32353 | -33.9798 | 0.89586
0.32963 | -33.9807 | 0.8672
0.33573 | -33.9819 | 0.84778
0.34184 | -33.9834 | 0.82714
0.34794 | -33.9857 | 0.80552
0.35405 | -33.988 | 0.78303
0.36015 | -33.9907 | 0.7595
0.36626 | -33.9928 | 0.73398
0.37236 | -33.9949 | 0.69888
0.37846 | -33.9971 | 0.65573
0.38457 | -33.9979 | 0.59777
0.39067 | -33.9988 | 0.55101
0.39678 | -33.9993 | 0.49815
0.40288 | -33.9996 | 0.44272
0.40899 | -33.9993 | 0.38907
0.41509 | -33.9984 | 0.33583
0.42119 | -33.9983 | 0.28984
0.4273 | -33.9981 | 0.24124
0.4334 | -33.9979 | 0.19268
0.43951 | -33.9982 | 0.14697
0.44561 | -33.9991 | 0.10521
0.45172 | -34.0008 | 0.072466
0.45782 | -34.0017 | 0.033977
0.46392 | -34.003 | -0.0035968
0.47003 | -34.0042 | -0.037114
0.47613 | -34.0052 | -0.072435
0.48224 | -34.0059 | -0.1084
0.48834 | -34.0049 | -0.1443
0.49445 | -34.0047 | -0.1808
0.50055 | -34.0041 | -0.2139
0.50665 | -34.0033 | -0.2494
0.51276 | -34.0022 | -0.28469
0.51886 | -34.001 | -0.3196
0.52497 | -33.9996 | -0.34439
0.53107 | -33.9986 | -0.37465
0.53717 | -33.9978 | -0.40363
0.54328 | -33.9967 | -0.42685
0.54938 | -33.9963 | -0.44975
0.55549 | -33.9962 | -0.47094
0.56159 | -33.9972 | -0.48253
0.5677 | -33.9974 | -0.49708
0.5738 | -33.9984 | -0.51344
0.5799 | -33.9991 | -0.52842
0.58601 | -34.0001 | -0.53836
0.59211 | -34.0019 | -0.54394
0.59822 | -34.0032 | -0.55312
0.60432 | -34.0047 | -0.56486
0.61043 | -34.006 | -0.57227
0.61653 | -34.0072 | -0.58202
0.62263 | -34.0088 | -0.59078
0.62874 | -34.0095 | -0.60055
0.63484 | -34.0103 | -0.60901
0.64095 | -34.0108 | -0.61848
0.64705 | -34.0111 | -0.62743
0.65316 | -34.0113 | -0.6393
0.65926 | -34.0113 | -0.65472
0.66536 | -34.0101 | -0.67191
0.67147 | -34.0097 | -0.68469
0.67757 | -34.0093 | -0.69855
0.68368 | -34.0084 | -0.71734
0.68978 | -34.0076 | -0.73433
0.69589 | -34.0065 | -0.75339
0.70199 | -34.0055 | -0.77669
0.70809 | -34.0045 | -0.79601
0.7142 | -34.0038 | -0.81677
0.7203 | -34.0026 | -0.83623
0.72641 | -34.0018 | -0.85924
0.73251 | -34.0011 | -0.87959
0.73862 | -34.0014 | -0.90195
0.74472 | -34.001 | -0.92227
0.75082 | -34.0007 | -0.94489
0.75693 | -34.001 | -0.96723
0.76303 | -34.0015 | -0.98968
0.76914 | -34.0035 | -1.0105
0.77524 | -34.0045 | -1.0306
0.78135 | -34.0058 | -1.0511
0.78745 | -34.0073 | -1.0714
0.79355 | -34.0086 | -1.0916
0.79966 | -34.0104 | -1.1114
0.80576 | -34.0129 | -1.1313
0.81187 | -34.0146 | -1.1533
0.81797 | -34.0166 | -1.175
0.82408 | -34.0184 | -1.1959
0.83018 | -34.0203 | -1.2169
0.83628 | -34.0223 | -1.2358
0.84239 | -34.0236 | -1.2567
0.84849 | -34.0251 | -1.279
0.8546 | -34.0267 | -1.2992
0.8607 | -34.028 | -1.3205
0.86681 | -34.0292 | -1.3398
0.87291 | -34.0303 | -1.3591
0.87901 | -34.0307 | -1.3777
0.88512 | -34.0315 | -1.3962
0.89122 | -34.0323 | -1.4149
0.89733 | -34.0328 | -1.4308
0.90343 | -34.0332 | -1.449
0.90953 | -34.0334 | -1.4637
0.91564 | -34.0334 | -1.48
0.92174 | -34.0337 | -1.4953
0.92785 | -34.0337 | -1.5102
0.93395 | -34.0338 | -1.5244
0.94006 | -34.034 | -1.5374
0.94616 | -34.0337 | -1.5495
0.95226 | -34.0338 | -1.5624
0.95837 | -34.0338 | -1.5766
0.96447 | -34.034 | -1.5896
0.97058 | -34.0338 | -1.6033
0.97668 | -34.0336 | -1.617
0.98279 | -34.0332 | -1.6299
0.98889 | -34.033 | -1.6407
0.99499 | -34.0329 | -1.6543
1.0011 | -34.0327 | -1.665
1.0072 | -34.0323 | -1.6766
1.0133 | -34.0319 | -1.6874
1.0194 | -34.0314 | -1.7004
1.0255 | -34.031 | -1.7143
1.0316 | -34.0303 | -1.7262
1.0377 | -34.0298 | -1.74
1.0438 | -34.0292 | -1.7527
1.0499 | -34.0283 | -1.7679
1.056 | -34.0275 | -1.7823
1.0621 | -34.0268 | -1.7968
1.0682 | -34.0262 | -1.8091
1.0743 | -34.0254 | -1.8247
1.0805 | -34.0249 | -1.839
1.0866 | -34.0238 | -1.854
1.0927 | -34.0229 | -1.8685
1.0988 | -34.0221 | -1.8834
1.1049 | -34.0212 | -1.8975
1.111 | -34.0203 | -1.9129
1.1171 | -34.0195 | -1.9273
1.1232 | -34.0189 | -1.9426
1.1293 | -34.0184 | -1.9565
1.1354 | -34.0181 | -1.9715
1.1415 | -34.0176 | -1.9871
1.1476 | -34.0169 | -2.0026
1.1537 | -34.0163 | -2.0174
1.1598 | -34.0164 | -2.0322
1.1659 | -34.0161 | -2.0454
1.172 | -34.0158 | -2.0593
1.1781 | -34.0155 | -2.0732
1.1842 | -34.0153 | -2.0874
1.1903 | -34.0157 | -2.0984
1.1964 | -34.0155 | -2.1118
1.2025 | -34.0156 | -2.1257
1.2086 | -34.0159 | -2.138
1.2147 | -34.0159 | -2.1506
1.2209 | -34.0163 | -2.1617
1.227 | -34.017 | -2.1724
1.2331 | -34.0174 | -2.1844
1.2392 | -34.0178 | -2.1962
1.2453 | -34.0183 | -2.2071
1.2514 | -34.0187 | -2.2169
1.2575 | -34.0193 | -2.2278
1.2636 | -34.0202 | -2.2367
1.2697 | -34.0208 | -2.2459
1.2758 | -34.0215 | -2.257
1.2819 | -34.0223 | -2.2668
1.288 | -34.023 | -2.2776
1.2941 | -34.0238 | -2.2859
1.3002 | -34.0245 | -2.2938
1.3063 | -34.0251 | -2.3028
1.3124 | -34.0258 | -2.3117
1.3185 | -34.0266 | -2.3208
1.3246 | -34.0273 | -2.3298
1.3307 | -34.0282 | -2.3383
1.3368 | -34.029 | -2.3469
1.3429 | -34.0297 | -2.3558
1.349 | -34.0302 | -2.365
1.3551 | -34.0309 | -2.3743
1.3613 | -34.0315 | -2.3825
1.3674 | -34.0321 | -2.3925
1.3735 | -34.0326 | -2.4022
1.3796 | -34.033 | -2.4112
1.3857 | -34.0334 | -2.4204
1.3918 | -34.0339 | -2.4305
1.3979 | -34.0345 | -2.4395
1.404 | -34.035 | -2.4509
1.4101 | -34.0354 | -2.4612
1.4162 | -34.0358 | -2.4718
1.4223 | -34.0363 | -2.4833
1.4284 | -34.0368 | -2.4933
1.4345 | -34.037 | -2.5048
1.4406 | -34.0373 | -2.5184
1.4467 | -34.0378 | -2.5301
1.4528 | -34.0382 | -2.5415
1.4589 | -34.0387 | -2.5534
1.465 | -34.039 | -2.5643
1.4711 | -34.0395 | -2.5756
1.4772 | -34.04 | -2.5891
1.4833 | -34.0405 | -2.6023
1.4894 | -34.041 | -2.6146
1.4955 | -34.0415 | -2.6266
1.5016 | -34.0421 | -2.6392
1.5078 | -34.0425 | -2.6533
1.5139 | -34.0429 | -2.6665
1.52 | -34.0433 | -2.6792
1.5261 | -34.0437 | -2.6919
1.5322 | -34.044 | -2.7061
1.5383 | -34.0444 | -2.7188
1.5444 | -34.0449 | -2.7315
1.5505 | -34.0454 | -2.7452
1.5566 | -34.0459 | -2.7583
1.5627 | -34.0464 | -2.7723
1.5688 | -34.0471 | -2.7861
1.5749 | -34.0474 | -2.7988
1.581 | -34.0477 | -2.8123
1.5871 | -34.0481 | -2.8262
1.5932 | -34.0485 | -2.8392
1.5993 | -34.0489 | -2.8536
1.6054 | -34.0493 | -2.867
1.6115 | -34.0497 | -2.8813
1.6176 | -34.05 | -2.894
1.6237 | -34.0502 | -2.9078
1.6298 | -34.0505 | -2.9205
1.6359 | -34.0508 | -2.9334
1.642 | -34.0513 | -2.9454
1.6482 | -34.0515 | -2.9585
1.6543 | -34.0518 | -2.9724
1.6604 | -34.0521 | -2.9845
1.6665 | -34.0522 | -2.997
1.6726 | -34.0524 | -3.0097
1.6787 | -34.0528 | -3.0223
1.6848 | -34.0527 | -3.035
1.6909 | -34.0529 | -3.0479
1.697 | -34.0531 | -3.0602
1.7031 | -34.0531 | -3.0729
1.7092 | -34.0531 | -3.0851
1.7153 | -34.0532 | -3.0971
1.7214 | -34.053 | -3.1095
1.7275 | -34.0529 | -3.1223
1.7336 | -34.0528 | -3.1354
1.7397 | -34.053 | -3.1479
1.7458 | -34.053 | -3.1602
1.7519 | -34.053 | -3.1725
1.758 | -34.053 | -3.1846
1.7641 | -34.0531 | -3.1963
1.7702 | -34.053 | -3.2095
1.7763 | -34.053 | -3.221
1.7824 | -34.0529 | -3.2336
1.7885 | -34.0529 | -3.2464
1.7947 | -34.053 | -3.2585
1.8008 | -34.053 | -3.2699
1.8069 | -34.0529 | -3.2816
1.813 | -34.053 | -3.2939
1.8191 | -34.0529 | -3.3069
1.8252 | -34.053 | -3.3199
1.8313 | -34.0531 | -3.3324
1.8374 | -34.0533 | -3.3444
1.8435 | -34.0533 | -3.3565
1.8496 | -34.0533 | -3.3699
1.8557 | -34.0535 | -3.3829
1.8618 | -34.0538 | -3.3954
1.8679 | -34.0539 | -3.4077
1.874 | -34.054 | -3.4199
1.8801 | -34.0541 | -3.4326
1.8862 | -34.0543 | -3.4447
1.8923 | -34.0545 | -3.4572
1.8984 | -34.0546 | -3.4695
1.9045 | -34.0546 | -3.4813
1.9106 | -34.0546 | -3.4938
1.9167 | -34.0547 | -3.5063
1.9228 | -34.0548 | -3.5184
1.9289 | -34.0549 | -3.5306
1.9351 | -34.0551 | -3.5416
1.9412 | -34.0553 | -3.5531
1.9473 | -34.0556 | -3.5654
1.9534 | -34.0557 | -3.5771
1.9595 | -34.0559 | -3.5892
1.9656 | -34.0562 | -3.6018
1.9717 | -34.0564 | -3.6138
1.9778 | -34.0567 | -3.6254
1.9839 | -34.0569 | -3.6372
1.99 | -34.0571 | -3.6485
1.9961 | -34.0572 | -3.661
2.0022 | -34.0574 | -3.6726
2.0083 | -34.0575 | -3.6841
2.0144 | -34.0577 | -3.6957
2.0205 | -34.058 | -3.7077
2.0266 | -34.0583 | -3.7194
2.0327 | -34.0586 | -3.7312
2.0388 | -34.059 | -3.743
2.0449 | -34.0591 | -3.7547
2.051 | -34.0594 | -3.7663
2.0571 | -34.0597 | -3.7779
2.0632 | -34.0599 | -3.7896
2.0693 | -34.0601 | -3.802
2.0754 | -34.0604 | -3.8136
2.0816 | -34.0607 | -3.826
2.0877 | -34.0608 | -3.8378
2.0938 | -34.0612 | -3.8502
2.0999 | -34.0614 | -3.8621
2.106 | -34.0616 | -3.874
2.1121 | -34.0617 | -3.8862
2.1182 | -34.062 | -3.8971
2.1243 | -34.0622 | -3.9081
2.1304 | -34.0624 | -3.9194
2.1365 | -34.0627 | -3.9314
2.1426 | -34.0629 | -3.9436
2.1487 | -34.0632 | -3.9551
2.1548 | -34.0637 | -3.9671
2.1609 | -34.0641 | -3.9788
2.167 | -34.0645 | -3.9902
2.1731 | -34.0648 | -4.0024
2.1792 | -34.065 | -4.0136
2.1853 | -34.0654 | -4.0258
2.1914 | -34.0657 | -4.0374
2.1975 | -34.0661 | -4.0492
2.2036 | -34.0663 | -4.0615
2.2097 | -34.0665 | -4.0732
2.2158 | -34.0669 | -4.084
2.222 | -34.0672 | -4.0946
2.2281 | -34.0674 | -4.1062
2.2342 | -34.0677 | -4.1178
2.2403 | -34.0679 | -4.1286
2.2464 | -34.0682 | -4.1395
2.2525 | -34.0685 | -4.1511
2.2586 | -34.0688 | -4.1618
2.2647 | -34.0691 | -4.1723
2.2708 | -34.0694 | -4.1838
2.2769 | -34.0698 | -4.1946
2.283 | -34.07 | -4.2059
2.2891 | -34.0702 | -4.217
2.2952 | -34.0703 | -4.2277
2.3013 | -34.0706 | -4.2387
2.3074 | -34.0709 | -4.2492
2.3135 | -34.0711 | -4.2608
2.3196 | -34.0713 | -4.2711
2.3257 | -34.0715 | -4.2819
2.3318 | -34.0716 | -4.2932
2.3379 | -34.0718 | -4.3041
2.344 | -34.0719 | -4.3157
2.3501 | -34.0722 | -4.3265
2.3562 | -34.0723 | -4.337
2.3623 | -34.0724 | -4.3475
2.3685 | -34.0725 | -4.3581
2.3746 | -34.0727 | -4.3695
2.3807 | -34.0729 | -4.3801
2.3868 | -34.0731 | -4.391
2.3929 | -34.0733 | -4.4018
2.399 | -34.0733 | -4.4124
2.4051 | -34.0736 | -4.423
2.4112 | -34.0738 | -4.4338
2.4173 | -34.074 | -4.4445
2.4234 | -34.074 | -4.4555
2.4295 | -34.0742 | -4.4664
2.4356 | -34.0743 | -4.4774
2.4417 | -34.0745 | -4.4887
2.4478 | -34.0746 | -4.4997
2.4539 | -34.0747 | -4.51
2.46 | -34.0748 | -4.5216
2.4661 | -34.075 | -4.5328
2.4722 | -34.0751 | -4.5431
2.4783 | -34.0752 | -4.5536
2.4844 | -34.0753 | -4.5647
2.4905 | -34.0753 | -4.5758
2.4966 | -34.0755 | -4.5867
2.5027 | -34.0756 | -4.5973
2.5089 | -34.0757 | -4.6083
2.515 | -34.0757 | -4.6193
2.5211 | -34.0758 | -4.6301
2.5272 | -34.0759 | -4.6411
2.5333 | -34.076 | -4.6522
2.5394 | -34.0762 | -4.6631
2.5455 | -34.0762 | -4.6747
2.5516 | -34.0764 | -4.6851
2.5577 | -34.0766 | -4.6967
2.5638 | -34.0767 | -4.7085
2.5699 | -34.0767 | -4.7201
2.576 | -34.0768 | -4.7318
2.5821 | -34.0769 | -4.7432
2.5882 | -34.077 | -4.7545
2.5943 | -34.0771 | -4.7654
2.6004 | -34.0773 | -4.7765
2.6065 | -34.0774 | -4.787
2.6126 | -34.0775 | -4.7975
2.6187 | -34.0776 | -4.8081
2.6248 | -34.0777 | -4.819
2.6309 | -34.0779 | -4.8306
2.637 | -34.0781 | -4.8424
2.6431 | -34.0783 | -4.8529
2.6492 | -34.0783 | -4.8644
2.6554 | -34.0785 | -4.8754
2.6615 | -34.0786 | -4.8865
2.6676 | -34.0789 | -4.8981
2.6737 | -34.0791 | -4.9094
2.6798 | -34.0794 | -4.9206
2.6859 | -34.0796 | -4.932
2.692 | -34.0798 | -4.943
2.6981 | -34.0799 | -4.9543
2.7042 | -34.0801 | -4.9656
2.7103 | -34.0803 | -4.9763
2.7164 | -34.0806 | -4.9874
2.7225 | -34.0809 | -4.9989
2.7286 | -34.0811 | -5.0101
2.7347 | -34.0813 | -5.0207
2.7408 | -34.0816 | -5.0309
2.7469 | -34.0818 | -5.0424
2.753 | -34.0821 | -5.0531
2.7591 | -34.0824 | -5.0645
2.7652 | -34.0826 | -5.0761
2.7713 | -34.0828 | -5.087
2.7774 | -34.083 | -5.0981
2.7835 | -34.0832 | -5.1093
2.7896 | -34.0834 | -5.1199
2.7958 | -34.0837 | -5.1308
2.8019 | -34.084 | -5.1413
2.808 | -34.0843 | -5.1522
2.8141 | -34.0846 | -5.1635
2.8202 | -34.0849 | -5.1745
2.8263 | -34.0852 | -5.1855
2.8324 | -34.0856 | -5.1971
2.8385 | -34.086 | -5.2081
2.8446 | -34.0863 | -5.219
2.8507 | -34.0867 | -5.23
2.8568 | -34.0871 | -5.2405
2.8629 | -34.0873 | -5.2513
2.869 | -34.0877 | -5.2626
2.8751 | -34.088 | -5.2735
2.8812 | -34.0885 | -5.2849
2.8873 | -34.0888 | -5.296
2.8934 | -34.0891 | -5.3072
2.8995 | -34.0894 | -5.3185
2.9056 | -34.0897 | -5.3298
2.9117 | -34.0901 | -5.341
2.9178 | -34.0905 | -5.3519
2.9239 | -34.0909 | -5.3636
2.93 | -34.0913 | -5.3753
2.9361 | -34.0916 | -5.3862
2.9423 | -34.092 | -5.3971
2.9484 | -34.0924 | -5.4083
2.9545 | -34.0928 | -5.4195
2.9606 | -34.0932 | -5.4311
2.9667 | -34.0936 | -5.4424
2.9728 | -34.0939 | -5.4539
2.9789 | -34.0944 | -5.4652
2.985 | -34.0949 | -5.4767
2.9911 | -34.0953 | -5.4882
2.9972 | -34.0957 | -5.4994
3.0033 | -34.0962 | -5.5112
3.0094 | -34.0966 | -5.5224
3.0155 | -34.097 | -5.5336
3.0216 | -34.0974 | -5.545
3.0277 | -34.0977 | -5.5569
3.0338 | -34.0982 | -5.5685
3.0399 | -34.0986 | -5.58
3.046 | -34.0989 | -5.5918
3.0521 | -34.0993 | -5.6031
3.0582 | -34.0996 | -5.6154
3.0643 | -34.0999 | -5.6271
3.0704 | -34.1003 | -5.6391
3.0765 | -34.1006 | -5.6504
3.0827 | -34.1009 | -5.6618
3.0888 | -34.1013 | -5.6733
3.0949 | -34.1017 | -5.6849
3.101 | -34.102 | -5.6965
3.1071 | -34.1024 | -5.7085
3.1132 | -34.1028 | -5.7205
3.1193 | -34.1031 | -5.7322
3.1254 | -34.1036 | -5.7437
3.1315 | -34.1039 | -5.7559
3.1376 | -34.1042 | -5.7675
3.1437 | -34.1046 | -5.7793
3.1498 | -34.105 | -5.7916
3.1559 | -34.1053 | -5.8032
3.162 | -34.1056 | -5.8155
3.1681 | -34.1058 | -5.8277
3.1742 | -34.1062 | -5.8395
3.1803 | -34.1065 | -5.8514
3.1864 | -34.1069 | -5.8629
3.1925 | -34.1072 | -5.8746
3.1986 | -34.1075 | -5.8868
3.2047 | -34.1078 | -5.8991
3.2108 | -34.1081 | -5.9104
3.2169 | -34.1084 | -5.9224
3.223 | -34.1086 | -5.9343
3.2292 | -34.1089 | -5.9464
3.2353 | -34.1091 | -5.9574
3.2414 | -34.1093 | -5.969
3.2475 | -34.1096 | -5.9806
3.2536 | -34.1099 | -5.9923
3.2597 | -34.1101 | -6.0034
3.2658 | -34.1104 | -6.0145
3.2719 | -34.1106 | -6.0257
3.278 | -34.1109 | -6.0376
3.2841 | -34.1111 | -6.0489
3.2902 | -34.1114 | -6.0605
3.2963 | -34.1116 | -6.0722
3.3024 | -34.1117 | -6.0835
3.3085 | -34.112 | -6.0953
3.3146 | -34.1122 | -6.1066
3.3207 | -34.1125 | -6.1183
3.3268 | -34.1127 | -6.1299
3.3329 | -34.1129 | -6.1409
3.339 | -34.1131 | -6.1518
3.3451 | -34.1132 | -6.1633
3.3512 | -34.1134 | -6.1741
3.3573 | -34.1137 | -6.1856
3.3634 | -34.1138 | -6.1969
3.3696 | -34.1141 | -6.2084
3.3757 | -34.1142 | -6.2194
3.3818 | -34.1144 | -6.2304
3.3879 | -34.1146 | -6.2417
3.394 | -34.1148 | -6.2533
3.4001 | -34.115 | -6.2643
3.4062 | -34.115 | -6.2754
3.4123 | -34.1152 | -6.2861
3.4184 | -34.1154 | -6.2976
3.4245 | -34.1156 | -6.3086
3.4306 | -34.1157 | -6.3194
3.4367 | -34.1159 | -6.3301
3.4428 | -34.1161 | -6.3407
3.4489 | -34.1162 | -6.3518
3.455 | -34.1163 | -6.3628
3.4611 | -34.1165 | -6.3743
3.4672 | -34.1167 | -6.3855
3.4733 | -34.1169 | -6.3967
3.4794 | -34.117 | -6.4077
3.4855 | -34.1172 | -6.4183
3.4916 | -34.1174 | -6.4288
3.4977 | -34.1175 | -6.4403
3.5038 | -34.1177 | -6.4514
3.5099 | -34.1178 | -6.4618
3.5161 | -34.1179 | -6.4724
3.5222 | -34.118 | -6.4829
3.5283 | -34.1181 | -6.4939
3.5344 | -34.1183 | -6.5048
3.5405 | -34.1184 | -6.5156
3.5466 | -34.1186 | -6.5267
3.5527 | -34.1187 | -6.5376
3.5588 | -34.1188 | -6.5484
3.5649 | -34.1189 | -6.5592
3.571 | -34.119 | -6.57
3.5771 | -34.1191 | -6.5813
3.5832 | -34.1193 | -6.5914
3.5893 | -34.1195 | -6.6022
3.5954 | -34.1197 | -6.6132
3.6015 | -34.1198 | -6.6243
3.6076 | -34.12 | -6.635
3.6137 | -34.1202 | -6.6452
3.6198 | -34.1205 | -6.6556
3.6259 | -34.1206 | -6.6667
3.632 | -34.1208 | -6.6775
3.6381 | -34.1209 | -6.6885
3.6442 | -34.1211 | -6.6992
3.6503 | -34.1213 | -6.7096
3.6565 | -34.1215 | -6.7205
3.6626 | -34.1216 | -6.7313
3.6687 | -34.1219 | -6.7422
3.6748 | -34.1221 | -6.7528
3.6809 | -34.1224 | -6.7636
3.687 | -34.1225 | -6.7748
3.6931 | -34.1228 | -6.7852
3.6992 | -34.123 | -6.796
3.7053 | -34.1232 | -6.8067
3.7114 | -34.1234 | -6.8176
3.7175 | -34.1236 | -6.8284
3.7236 | -34.1239 | -6.8386
3.7297 | -34.1241 | -6.8496
3.7358 | -34.1243 | -6.8607
3.7419 | -34.1245 | -6.8718
3.748 | -34.1247 | -6.8825
3.7541 | -34.1249 | -6.8933
3.7602 | -34.1252 | -6.9039
3.7663 | -34.1254 | -6.9146
3.7724 | -34.1257 | -6.9254
3.7785 | -34.126 | -6.936
3.7846 | -34.1262 | -6.9466
3.7907 | -34.1265 | -6.9578
3.7969 | -34.1267 | -6.9687
3.803 | -34.127 | -6.9791
3.8091 | -34.1272 | -6.9898
3.8152 | -34.1275 | -7.0005
3.8213 | -34.1277 | -7.0115
3.8274 | -34.128 | -7.022
3.8335 | -34.1284 | -7.0329
3.8396 | -34.1287 | -7.0437
3.8457 | -34.129 | -7.0546
3.8518 | -34.1293 | -7.0654
3.8579 | -34.1297 | -7.0761
3.864 | -34.13 | -7.0871
3.8701 | -34.1303 | -7.0983
3.8762 | -34.1306 | -7.1089
3.8823 | -34.1308 | -7.1199
3.8884 | -34.1311 | -7.1308
3.8945 | -34.1315 | -7.1419
3.9006 | -34.1318 | -7.1531
3.9067 | -34.1321 | -7.1644
3.9128 | -34.1324 | -7.1753
3.9189 | -34.1327 | -7.1859
3.925 | -34.133 | -7.1969
3.9311 | -34.1333 | -7.2073
3.9372 | -34.1336 | -7.2184
3.9434 | -34.1339 | -7.2293
3.9495 | -34.1343 | -7.2403
3.9556 | -34.1346 | -7.2509
3.9617 | -34.1349 | -7.2614
3.9678 | -34.1353 | -7.2719
3.9739 | -34.1356 | -7.2828
3.98 | -34.1358 | -7.2938
3.9861 | -34.1361 | -7.3045
3.9922 | -34.1364 | -7.3154
3.9983 | -34.1367 | -7.3258
4.0044 | -34.137 | -7.3362
4.0105 | -34.1373 | -7.3471
4.0166 | -34.1377 | -7.358
4.0227 | -34.138 | -7.3688
4.0288 | -34.1382 | -7.3801
4.0349 | -34.1385 | -7.3907
4.041 | -34.1388 | -7.4014
4.0471 | -34.139 | -7.4124
4.0532 | -34.1393 | -7.4228
4.0593 | -34.1397 | -7.4338
4.0654 | -34.14 | -7.4444
4.0715 | -34.1403 | -7.4557
4.0776 | -34.1406 | -7.4664
4.0838 | -34.1409 | -7.4777
4.0899 | -34.1412 | -7.4882
4.096 | -34.1415 | -7.4989
4.1021 | -34.1418 | -7.5095
4.1082 | -34.142 | -7.52
4.1143 | -34.1423 | -7.5305
4.1204 | -34.1426 | -7.5409
4.1265 | -34.1429 | -7.5514
4.1326 | -34.1432 | -7.5618
4.1387 | -34.1436 | -7.5724
4.1448 | -34.1438 | -7.5832
4.1509 | -34.1441 | -7.5939
4.157 | -34.1443 | -7.6051
4.1631 | -34.1446 | -7.616
4.1692 | -34.1449 | -7.6266
4.1753 | -34.1452 | -7.6372
4.1814 | -34.1454 | -7.6475
4.1875 | -34.1457 | -7.6585
4.1936 | -34.146 | -7.6694
4.1997 | -34.1462 | -7.6797
4.2058 | -34.1465 | -7.6907
4.2119 | -34.1468 | -7.7012
4.218 | -34.147 | -7.7116
4.2241 | -34.1472 | -7.722
4.2303 | -34.1476 | -7.7325
4.2364 | -34.1478 | -7.7431
4.2425 | -34.148 | -7.7537
4.2486 | -34.1482 | -7.7646
4.2547 | -34.1484 | -7.7749
4.2608 | -34.1486 | -7.7857
4.2669 | -34.1488 | -7.7959
4.273 | -34.1491 | -7.8067
4.2791 | -34.1493 | -7.8173
4.2852 | -34.1496 | -7.8281
4.2913 | -34.1497 | -7.8384
4.2974 | -34.15 | -7.8493
4.3035 | -34.1502 | -7.8602
4.3096 | -34.1504 | -7.871
4.3157 | -34.1507 | -7.8813
4.3218 | -34.1509 | -7.8918
4.3279 | -34.1512 | -7.9024
4.334 | -34.1513 | -7.913
4.3401 | -34.1516 | -7.9231
4.3462 | -34.1519 | -7.9337
4.3523 | -34.1521 | -7.9439
4.3584 | -34.1523 | -7.954
4.3645 | -34.1525 | -7.9644
4.3707 | -34.1527 | -7.9746
4.3768 | -34.153 | -7.9851
4.3829 | -34.1532 | -7.996
4.389 | -34.1534 | -8.0064
4.3951 | -34.1536 | -8.0168
4.4012 | -34.1538 | -8.0276
4.4073 | -34.154 | -8.0379
4.4134 | -34.1542 | -8.0483
4.4195 | -34.1545 | -8.0587
4.4256 | -34.1548 | -8.069
4.4317 | -34.155 | -8.0794
4.4378 | -34.1552 | -8.0895
4.4439 | -34.1554 | -8.0996
4.45 | -34.1556 | -8.11
4.4561 | -34.1557 | -8.1205
4.4622 | -34.1559 | -8.131
4.4683 | -34.1562 | -8.1414
4.4744 | -34.1564 | -8.1521
4.4805 | -34.1566 | -8.1621
4.4866 | -34.1568 | -8.1725
4.4927 | -34.157 | -8.1832
4.4988 | -34.1573 | -8.1936
4.5049 | -34.1575 | -8.2035
4.511 | -34.1578 | -8.2136
4.5172 | -34.158 | -8.224
4.5233 | -34.1581 | -8.2348
4.5294 | -34.1584 | -8.245
4.5355 | -34.1586 | -8.2554
4.5416 | -34.1588 | -8.2658
4.5477 | -34.159 | -8.2762
4.5538 | -34.1592 | -8.2861
4.5599 | -34.1595 | -8.2964
4.566 | -34.1597 | -8.3067
4.5721 | -34.1599 | -8.3171
4.5782 | -34.1602 | -8.3278
4.5843 | -34.1604 | -8.3382
4.5904 | -34.1607 | -8.3487
4.5965 | -34.1609 | -8.3589
4.6026 | -34.1611 | -8.3698
4.6087 | -34.1613 | -8.3801
4.6148 | -34.1615 | -8.3908
4.6209 | -34.1618 | -8.4011
4.627 | -34.1621 | -8.4117
4.6331 | -34.1622 | -8.422
4.6392 | -34.1624 | -8.4325
4.6453 | -34.1627 | -8.4428
4.6514 | -34.1629 | -8.4533
4.6576 | -34.1631 | -8.4641
4.6637 | -34.1633 | -8.4741
4.6698 | -34.1635 | -8.4843
4.6759 | -34.1637 | -8.4951
4.682 | -34.1639 | -8.5054
4.6881 | -34.1641 | -8.5157
4.6942 | -34.1643 | -8.5261
4.7003 | -34.1645 | -8.5365
4.7064 | -34.1648 | -8.5467
4.7125 | -34.1651 | -8.5574
4.7186 | -34.1652 | -8.5676
4.7247 | -34.1655 | -8.5781
4.7308 | -34.1658 | -8.5884
4.7369 | -34.166 | -8.5989
4.743 | -34.1662 | -8.6097
4.7491 | -34.1664 | -8.6204
4.7552 | -34.1666 | -8.631
4.7613 | -34.1668 | -8.6415
4.7674 | -34.167 | -8.6518
4.7735 | -34.1673 | -8.6627
4.7796 | -34.1675 | -8.6734
4.7857 | -34.1678 | -8.684
4.7918 | -34.168 | -8.6943
4.7979 | -34.1683 | -8.7046
4.8041 | -34.1685 | -8.7151
4.8102 | -34.1688 | -8.7252
4.8163 | -34.169 | -8.7356
4.8224 | -34.1692 | -8.7461
4.8285 | -34.1695 | -8.7562
4.8346 | -34.1698 | -8.7665
4.8407 | -34.1699 | -8.777
4.8468 | -34.1702 | -8.7878
4.8529 | -34.1705 | -8.7985
4.859 | -34.1707 | -8.8093
4.8651 | -34.171 | -8.8195
4.8712 | -34.1712 | -8.8298
4.8773 | -34.1714 | -8.8403
4.8834 | -34.1717 | -8.8508
4.8895 | -34.1719 | -8.8611
4.8956 | -34.1721 | -8.8715
4.9017 | -34.1723 | -8.8818
4.9078 | -34.1725 | -8.8926
4.9139 | -34.1728 | -8.9026
4.92 | -34.1731 | -8.9127
4.9261 | -34.1734 | -8.9228
4.9322 | -34.1736 | -8.9333
4.9383 | -34.1739 | -8.9437
4.9445 | -34.1742 | -8.9542
4.9506 | -34.1744 | -8.9644
4.9567 | -34.1746 | -8.9746
4.9628 | -34.1749 | -8.985
4.9689 | -34.1751 | -8.9958
4.975 | -34.1753 | -9.0066
4.9811 | -34.1756 | -9.0168
4.9872 | -34.1759 | -9.0271
4.9933 | -34.1762 | -9.0378
4.9994 | -34.1765 | -9.0487
5.0055 | -34.1768 | -9.0592
5.0116 | -34.1771 | -9.0701
5.0177 | -34.1774 | -9.0808
5.0238 | -34.1777 | -9.0916
5.0299 | -34.178 | -9.1025
5.036 | -34.1783 | -9.113
5.0421 | -34.1785 | -9.1233
5.0482 | -34.1788 | -9.1339
5.0543 | -34.1791 | -9.1441
5.0604 | -34.1794 | -9.1545
5.0665 | -34.1797 | -9.1648
5.0726 | -34.1799 | -9.1756
5.0787 | -34.1802 | -9.1863
5.0848 | -34.1805 | -9.1969
5.091 | -34.1808 | -9.2072
5.0971 | -34.1811 | -9.2173
5.1032 | -34.1814 | -9.2282
5.1093 | -34.1817 | -9.2387
5.1154 | -34.1819 | -9.2489
5.1215 | -34.1822 | -9.2592
5.1276 | -34.1825 | -9.2695
5.1337 | -34.1828 | -9.2797
5.1398 | -34.1831 | -9.2899
5.1459 | -34.1834 | -9.3007
5.152 | -34.1836 | -9.3113
5.1581 | -34.1839 | -9.322
5.1642 | -34.1842 | -9.3326
5.1703 | -34.1845 | -9.3431
5.1764 | -34.1848 | -9.3534
5.1825 | -34.1851 | -9.3637
5.1886 | -34.1854 | -9.3738
5.1947 | -34.1857 | -9.3844
5.2008 | -34.186 | -9.395
5.2069 | -34.1863 | -9.4058
5.213 | -34.1866 | -9.416
5.2191 | -34.1869 | -9.4266
5.2252 | -34.1872 | -9.4371
5.2314 | -34.1874 | -9.4478
5.2375 | -34.1877 | -9.4581
5.2436 | -34.188 | -9.4687
5.2497 | -34.1883 | -9.4789
5.2558 | -34.1887 | -9.4897
5.2619 | -34.1889 | -9.5005
5.268 | -34.1893 | -9.5112
5.2741 | -34.1896 | -9.5218
5.2802 | -34.1899 | -9.5328
5.2863 | -34.1902 | -9.5432
5.2924 | -34.1905 | -9.5539
5.2985 | -34.1909 | -9.5645
5.3046 | -34.1912 | -9.5751
5.3107 | -34.1915 | -9.5859
5.3168 | -34.1918 | -9.5965
5.3229 | -34.1921 | -9.6071
5.329 | -34.1924 | -9.6177
5.3351 | -34.1927 | -9.6279
5.3412 | -34.193 | -9.6384
5.3473 | -34.1933 | -9.6491
5.3534 | -34.1937 | -9.6598
5.3595 | -34.1939 | -9.6704
5.3656 | -34.1943 | -9.6809
5.3717 | -34.1946 | -9.6919
5.3779 | -34.1949 | -9.7028
5.384 | -34.1952 | -9.7134
5.3901 | -34.1955 | -9.7242
5.3962 | -34.1958 | -9.7346
5.4023 | -34.1962 | -9.7455
5.4084 | -34.1965 | -9.756
5.4145 | -34.1969 | -9.7663
5.4206 | -34.1971 | -9.777
5.4267 | -34.1975 | -9.7879
5.4328 | -34.1978 | -9.7986
5.4389 | -34.1981 | -9.8089
5.445 | -34.1984 | -9.8199
5.4511 | -34.1987 | -9.8302
5.4572 | -34.1991 | -9.841
5.4633 | -34.1994 | -9.8516
5.4694 | -34.1996 | -9.8622
5.4755 | -34.2 | -9.8726
5.4816 | -34.2003 | -9.8832
5.4877 | -34.2006 | -9.8937
5.4938 | -34.2009 | -9.904
5.4999 | -34.2013 | -9.9145
5.506 | -34.2016 | -9.9251
5.5121 | -34.2019 | -9.9356
5.5183 | -34.2023 | -9.946
5.5244 | -34.2025 | -9.9569
5.5305 | -34.2029 | -9.9674
5.5366 | -34.2032 | -9.9777
5.5427 | -34.2035 | -9.9885
5.5488 | -34.2038 | -9.9991
5.5549 | -34.2041 | -10.0098
5.561 | -34.2044 | -10.0206
5.5671 | -34.2048 | -10.0311
5.5732 | -34.2051 | -10.0419
5.5793 | -34.2054 | -10.0529
5.5854 | -34.2058 | -10.0635
5.5915 | -34.2061 | -10.0744
5.5976 | -34.2064 | -10.0851
5.6037 | -34.2067 | -10.0958
5.6098 | -34.207 | -10.1064
5.6159 | -34.2073 | -10.1168
5.622 | -34.2076 | -10.1276
5.6281 | -34.208 | -10.1383
5.6342 | -34.2083 | -10.1493
5.6403 | -34.2086 | -10.1599
5.6464 | -34.2089 | -10.1707
5.6525 | -34.2093 | -10.1815
5.6586 | -34.2096 | -10.1918
5.6648 | -34.21 | -10.2025
5.6709 | -34.2103 | -10.2133
5.677 | -34.2106 | -10.224
5.6831 | -34.211 | -10.2347
5.6892 | -34.2113 | -10.2456
5.6953 | -34.2116 | -10.2562
5.7014 | -34.2119 | -10.267
5.7075 | -34.2122 | -10.2776
5.7136 | -34.2125 | -10.2882
5.7197 | -34.2128 | -10.2988
5.7258 | -34.2131 | -10.3092
5.7319 | -34.2135 | -10.32
5.738 | -34.2138 | -10.3306
5.7441 | -34.2142 | -10.3412
5.7502 | -34.2145 | -10.3515
5.7563 | -34.2149 | -10.3622
5.7624 | -34.2152 | -10.3731
5.7685 | -34.2155 | -10.3838
5.7746 | -34.2158 | -10.3947
5.7807 | -34.2161 | -10.4054
5.7868 | -34.2164 | -10.4163
5.7929 | -34.2167 | -10.4269
5.799 | -34.217 | -10.4374
5.8052 | -34.2174 | -10.4479
5.8113 | -34.2177 | -10.4584
5.8174 | -34.218 | -10.4689
5.8235 | -34.2183 | -10.4795
5.8296 | -34.2186 | -10.4904
5.8357 | -34.2189 | -10.501
5.8418 | -34.2192 | -10.5118
5.8479 | -34.2195 | -10.5225
5.854 | -34.2199 | -10.5333
5.8601 | -34.2202 | -10.5438
5.8662 | -34.2205 | -10.5545
5.8723 | -34.2208 | -10.565
5.8784 | -34.2211 | -10.5756
5.8845 | -34.2214 | -10.5861
5.8906 | -34.2217 | -10.5968
5.8967 | -34.222 | -10.6072
5.9028 | -34.2224 | -10.6176
5.9089 | -34.2228 | -10.6281
5.915 | -34.2231 | -10.6385
5.9211 | -34.2234 | -10.6491
5.9272 | -34.2237 | -10.6594
5.9333 | -34.224 | -10.6699
5.9394 | -34.2244 | -10.6805
5.9455 | -34.2247 | -10.6909
5.9517 | -34.2251 | -10.7013
5.9578 | -34.2254 | -10.712
5.9639 | -34.2258 | -10.7222
5.97 | -34.2261 | -10.7329
5.9761 | -34.2265 | -10.7434
5.9822 | -34.2268 | -10.754
5.9883 | -34.2271 | -10.7646
5.9944 | -34.2274 | -10.7753
6.0005 | -34.2278 | -10.786
6.0066 | -34.2281 | -10.7964
6.0127 | -34.2285 | -10.8067
6.0188 | -34.2288 | -10.8175
6.0249 | -34.2291 | -10.8282
6.031 | -34.2294 | -10.8386
6.0371 | -34.2297 | -10.8492
6.0432 | -34.2301 | -10.8596
6.0493 | -34.2304 | -10.8699
6.0554 | -34.2308 | -10.8803
6.0615 | -34.2311 | -10.8909
6.0676 | -34.2314 | -10.9016
6.0737 | -34.2318 | -10.9123
6.0798 | -34.2321 | -10.9228
6.0859 | -34.2325 | -10.9333
6.0921 | -34.2328 | -10.9435
6.0982 | -34.2331 | -10.9541
6.1043 | -34.2335 | -10.9645
6.1104 | -34.2338 | -10.9752
6.1165 | -34.2342 | -10.9852
6.1226 | -34.2345 | -10.9957
6.1287 | -34.2348 | -11.0061
6.1348 | -34.2352 | -11.0166
6.1409 | -34.2355 | -11.027
6.147 | -34.2359 | -11.0374
6.1531 | -34.2362 | -11.0479
6.1592 | -34.2365 | -11.0582
6.1653 | -34.2368 | -11.0686
6.1714 | -34.2371 | -11.079
6.1775 | -34.2375 | -11.0897
6.1836 | -34.2378 | -11.1002
6.1897 | -34.2382 | -11.1106
6.1958 | -34.2385 | -11.1213
6.2019 | -34.2389 | -11.132
6.208 | -34.2392 | -11.1425
6.2141 | -34.2395 | -11.1528
6.2202 | -34.2399 | -11.1634
6.2263 | -34.2403 | -11.1738
6.2325 | -34.2406 | -11.1839
6.2386 | -34.2409 | -11.194
6.2447 | -34.2413 | -11.2047
6.2508 | -34.2416 | -11.2152
6.2569 | -34.242 | -11.2255
6.263 | -34.2423 | -11.236
6.2691 | -34.2426 | -11.2463
6.2752 | -34.2429 | -11.2568
6.2813 | -34.2432 | -11.2672
6.2874 | -34.2436 | -11.2775
6.2935 | -34.2439 | -11.2878
6.2996 | -34.2442 | -11.2982
6.3057 | -34.2446 | -11.3087
6.3118 | -34.2448 | -11.3193
6.3179 | -34.2452 | -11.3296
6.324 | -34.2455 | -11.34
6.3301 | -34.2459 | -11.3503
6.3362 | -34.2462 | -11.3606
6.3423 | -34.2466 | -11.3712
6.3484 | -34.2469 | -11.3817
6.3545 | -34.2472 | -11.3917
6.3606 | -34.2475 | -11.4018
6.3667 | -34.2478 | -11.412
6.3728 | -34.2482 | -11.4225
6.379 | -34.2485 | -11.4329
6.3851 | -34.2489 | -11.443
6.3912 | -34.2492 | -11.4533
6.3973 | -34.2495 | -11.4637
6.4034 | -34.2498 | -11.4738
6.4095 | -34.2501 | -11.4842
6.4156 | -34.2505 | -11.4946
6.4217 | -34.2508 | -11.5048
6.4278 | -34.2511 | -11.5153
6.4339 | -34.2515 | -11.5255
6.44 | -34.2518 | -11.5358
6.4461 | -34.2521 | -11.5464
6.4522 | -34.2525 | -11.5564
6.4583 | -34.2527 | -11.5669
6.4644 | -34.2531 | -11.5774
6.4705 | -34.2534 | -11.5873
6.4766 | -34.2537 | -11.5977
6.4827 | -34.2541 | -11.6081
6.4888 | -34.2544 | -11.6182
6.4949 | -34.2547 | -11.6282
6.501 | -34.255 | -11.6386
6.5071 | -34.2554 | -11.6487
6.5132 | -34.2557 | -11.659
6.5194 | -34.256 | -11.6694
6.5255 | -34.2563 | -11.6794
6.5316 | -34.2566 | -11.6899
6.5377 | -34.257 | -11.7001
6.5438 | -34.2573 | -11.7102
6.5499 | -34.2576 | -11.7203
6.556 | -34.2579 | -11.7306
6.5621 | -34.2582 | -11.7411
6.5682 | -34.2586 | -11.7514
6.5743 | -34.259 | -11.7616
6.5804 | -34.2593 | -11.772
6.5865 | -34.2596 | -11.7823
6.5926 | -34.2599 | -11.7925
6.5987 | -34.2603 | -11.8024
6.6048 | -34.2606 | -11.8127
6.6109 | -34.2609 | -11.8229
6.617 | -34.2612 | -11.8333
6.6231 | -34.2615 | -11.8435
6.6292 | -34.2619 | -11.8538
6.6353 | -34.2622 | -11.8643
6.6414 | -34.2626 | -11.8745
6.6475 | -34.2629 | -11.8847
6.6536 | -34.2632 | -11.8949
6.6597 | -34.2635 | -11.9048
6.6659 | -34.2638 | -11.915
6.672 | -34.2642 | -11.9252
6.6781 | -34.2645 | -11.9354
6.6842 | -34.2649 | -11.9455
6.6903 | -34.2652 | -11.9556
6.6964 | -34.2656 | -11.9659
6.7025 | -34.2659 | -11.9759
6.7086 | -34.2662 | -11.9861
6.7147 | -34.2665 | -11.9963
6.7208 | -34.2668 | -12.0067
6.7269 | -34.2672 | -12.0168
6.733 | -34.2675 | -12.0268
6.7391 | -34.2678 | -12.037
6.7452 | -34.2682 | -12.0469
6.7513 | -34.2685 | -12.0568
6.7574 | -34.2688 | -12.0671
6.7635 | -34.2691 | -12.0771
6.7696 | -34.2695 | -12.0871
6.7757 | -34.2699 | -12.0973
6.7818 | -34.2702 | -12.1075
6.7879 | -34.2706 | -12.1174
6.794 | -34.2709 | -12.1274
6.8001 | -34.2712 | -12.1375
6.8063 | -34.2715 | -12.1476
6.8124 | -34.2718 | -12.1579
6.8185 | -34.2721 | -12.1682
6.8246 | -34.2724 | -12.1782
6.8307 | -34.2727 | -12.1882
6.8368 | -34.2731 | -12.1986
6.8429 | -34.2734 | -12.2087
6.849 | -34.2738 | -12.2189
6.8551 | -34.2742 | -12.2289
6.8612 | -34.2745 | -12.239
6.8673 | -34.2748 | -12.249
6.8734 | -34.2751 | -12.2593
6.8795 | -34.2754 | -12.2695
6.8856 | -34.2758 | -12.2796
6.8917 | -34.2761 | -12.2896
6.8978 | -34.2764 | -12.2994
6.9039 | -34.2768 | -12.3092
6.91 | -34.2771 | -12.3194
6.9161 | -34.2775 | -12.3292
6.9222 | -34.2778 | -12.3393
6.9283 | -34.2782 | -12.3491
6.9344 | -34.2785 | -12.3593
6.9405 | -34.2789 | -12.3693
6.9466 | -34.2793 | -12.3796
6.9528 | -34.2796 | -12.3899
6.9589 | -34.28 | -12.4001
6.965 | -34.2803 | -12.4103
6.9711 | -34.2807 | -12.4205
6.9772 | -34.2811 | -12.4306
6.9833 | -34.2814 | -12.4406
6.9894 | -34.2817 | -12.4509
6.9955 | -34.2821 | -12.4607
7.0016 | -34.2824 | -12.4708
7.0077 | -34.2828 | -12.4808
7.0138 | -34.2831 | -12.4908
7.0199 | -34.2834 | -12.5008
7.026 | -34.2838 | -12.5108
7.0321 | -34.2841 | -12.5206
7.0382 | -34.2845 | -12.5303
7.0443 | -34.2849 | -12.5406
7.0504 | -34.2852 | -12.5507
7.0565 | -34.2855 | -12.5608
7.0626 | -34.2859 | -12.5706
7.0687 | -34.2862 | -12.5802
7.0748 | -34.2866 | -12.59
7.0809 | -34.287 | -12.5999
7.087 | -34.2873 | -12.61
7.0932 | -34.2876 | -12.6201
7.0993 | -34.288 | -12.6301
7.1054 | -34.2884 | -12.6399
7.1115 | -34.2887 | -12.6497
7.1176 | -34.2891 | -12.6597
7.1237 | -34.2894 | -12.6696
7.1298 | -34.2897 | -12.6795
7.1359 | -34.2901 | -12.6895
7.142 | -34.2904 | -12.6992
7.1481 | -34.2908 | -12.7091
7.1542 | -34.2911 | -12.7191
7.1603 | -34.2915 | -12.7291
7.1664 | -34.2919 | -12.739
7.1725 | -34.2922 | -12.7491
7.1786 | -34.2926 | -12.759
7.1847 | -34.2929 | -12.769
7.1908 | -34.2933 | -12.7792
7.1969 | -34.2936 | -12.7893
7.203 | -34.294 | -12.7991
7.2091 | -34.2943 | -12.809
7.2152 | -34.2947 | -12.819
7.2213 | -34.295 | -12.8288
7.2274 | -34.2954 | -12.8387
7.2335 | -34.2957 | -12.8486
7.2397 | -34.2961 | -12.8585
7.2458 | -34.2965 | -12.8682
7.2519 | -34.2969 | -12.8782
7.258 | -34.2972 | -12.8883
7.2641 | -34.2976 | -12.8982
7.2702 | -34.2979 | -12.9084
7.2763 | -34.2983 | -12.9184
7.2824 | -34.2987 | -12.9285
7.2885 | -34.299 | -12.9385
7.2946 | -34.2994 | -12.9484
7.3007 | -34.2997 | -12.9585
7.3068 | -34.3 | -12.9687
7.3129 | -34.3004 | -12.9788
7.319 | -34.3008 | -12.9888
7.3251 | -34.3012 | -12.9986
7.3312 | -34.3015 | -13.0085
7.3373 | -34.3019 | -13.0182
7.3434 | -34.3023 | -13.0281
7.3495 | -34.3027 | -13.0379
7.3556 | -34.303 | -13.048
7.3617 | -34.3034 | -13.0581
7.3678 | -34.3037 | -13.068
7.3739 | -34.3041 | -13.0782
7.3801 | -34.3044 | -13.0884
7.3862 | -34.3048 | -13.0983
7.3923 | -34.3052 | -13.1085
7.3984 | -34.3055 | -13.1182
7.4045 | -34.306 | -13.1281
7.4106 | -34.3064 | -13.138
7.4167 | -34.3067 | -13.148
7.4228 | -34.3071 | -13.1581
7.4289 | -34.3074 | -13.168
7.435 | -34.3078 | -13.1778
7.4411 | -34.3082 | -13.1879
7.4472 | -34.3085 | -13.1979
7.4533 | -34.3089 | -13.2076
7.4594 | -34.3093 | -13.2179
7.4655 | -34.3097 | -13.2276
7.4716 | -34.31 | -13.2377
7.4777 | -34.3104 | -13.2477
7.4838 | -34.3108 | -13.2575
7.4899 | -34.3111 | -13.2674
7.496 | -34.3115 | -13.2772
7.5021 | -34.3119 | -13.2871
7.5082 | -34.3122 | -13.297
7.5143 | -34.3126 | -13.307
7.5204 | -34.313 | -13.3171
7.5266 | -34.3133 | -13.3269
7.5327 | -34.3137 | -13.3368
7.5388 | -34.3141 | -13.3471
7.5449 | -34.3145 | -13.3573
7.551 | -34.3148 | -13.3671
7.5571 | -34.3152 | -13.3772
7.5632 | -34.3156 | -13.3872
7.5693 | -34.316 | -13.397
7.5754 | -34.3164 | -13.4068
7.5815 | -34.3167 | -13.4168
7.5876 | -34.3171 | -13.427
7.5937 | -34.3175 | -13.4372
7.5998 | -34.3179 | -13.4471
7.6059 | -34.3183 | -13.4572
7.612 | -34.3187 | -13.4672
7.6181 | -34.319 | -13.4771
7.6242 | -34.3194 | -13.4873
7.6303 | -34.3198 | -13.4976
7.6364 | -34.3201 | -13.5076
7.6425 | -34.3205 | -13.5176
7.6486 | -34.3209 | -13.5276
7.6547 | -34.3213 | -13.5375
7.6608 | -34.3216 | -13.5474
7.667 | -34.322 | -13.5576
7.6731 | -34.3224 | -13.5676
7.6792 | -34.3228 | -13.5779
7.6853 | -34.3231 | -13.588
7.6914 | -34.3235 | -13.5981
7.6975 | -34.3239 | -13.6083
7.7036 | -34.3243 | -13.6184
7.7097 | -34.3247 | -13.6285
7.7158 | -34.3251 | -13.6386
7.7219 | -34.3255 | -13.6485
7.728 | -34.3258 | -13.6586
7.7341 | -34.3262 | -13.6688
7.7402 | -34.3266 | -13.6789
7.7463 | -34.327 | -13.689
7.7524 | -34.3273 | -13.6993
7.7585 | -34.3277 | -13.7094
7.7646 | -34.3281 | -13.7193
7.7707 | -34.3285 | -13.7297
7.7768 | -34.3288 | -13.7399
7.7829 | -34.3292 | -13.7503
7.789 | -34.3296 | -13.7606
7.7951 | -34.3299 | -13.7706
7.8012 | -34.3303 | -13.7807
7.8073 | -34.3308 | -13.7911
7.8135 | -34.3311 | -13.8011
7.8196 | -34.3315 | -13.8115
7.8257 | -34.3319 | -13.8218
7.8318 | -34.3322 | -13.832
7.8379 | -34.3327 | -13.8419
7.844 | -34.333 | -13.852
7.8501 | -34.3334 | -13.8622
7.8562 | -34.3338 | -13.8725
7.8623 | -34.3342 | -13.8825
7.8684 | -34.3346 | -13.8927
7.8745 | -34.3349 | -13.9028
7.8806 | -34.3353 | -13.9132
7.8867 | -34.3357 | -13.9231
7.8928 | -34.3361 | -13.933
7.8989 | -34.3365 | -13.9431
7.905 | -34.3369 | -13.9533
7.9111 | -34.3373 | -13.9634
7.9172 | -34.3376 | -13.9735
7.9233 | -34.338 | -13.9837
7.9294 | -34.3384 | -13.994
7.9355 | -34.3388 | -14.0039
7.9416 | -34.3391 | -14.0141
7.9477 | -34.3395 | -14.024
7.9539 | -34.3399 | -14.034
7.96 | -34.3403 | -14.0443
7.9661 | -34.3407 | -14.0545
7.9722 | -34.3411 | -14.0645
7.9783 | -34.3415 | -14.0749
7.9844 | -34.3419 | -14.085
7.9905 | -34.3423 | -14.0953
7.9966 | -34.3426 | -14.1056
8.0027 | -34.343 | -14.1158
8.0088 | -34.3434 | -14.1259
8.0149 | -34.3437 | -14.1361
8.021 | -34.3441 | -14.1462
8.0271 | -34.3445 | -14.1567
8.0332 | -34.3449 | -14.1668
8.0393 | -34.3453 | -14.1771
8.0454 | -34.3457 | -14.1875
8.0515 | -34.3461 | -14.1976
8.0576 | -34.3465 | -14.208
8.0637 | -34.3468 | -14.2181
8.0698 | -34.3472 | -14.2281
8.0759 | -34.3476 | -14.2381
8.082 | -34.348 | -14.2481
8.0881 | -34.3484 | -14.2584
8.0942 | -34.3488 | -14.2686
8.1004 | -34.3492 | -14.2789
8.1065 | -34.3496 | -14.2889
8.1126 | -34.3499 | -14.2991
8.1187 | -34.3503 | -14.3094
8.1248 | -34.3507 | -14.3193
8.1309 | -34.3511 | -14.3293
8.137 | -34.3515 | -14.3393
8.1431 | -34.3519 | -14.3496
8.1492 | -34.3522 | -14.3597
8.1553 | -34.3526 | -14.3699
8.1614 | -34.353 | -14.3803
8.1675 | -34.3534 | -14.3904
8.1736 | -34.3537 | -14.4006
8.1797 | -34.3541 | -14.4109
8.1858 | -34.3545 | -14.4211
8.1919 | -34.3549 | -14.4313
8.198 | -34.3552 | -14.4412
8.2041 | -34.3557 | -14.4515
8.2102 | -34.356 | -14.4615
8.2163 | -34.3564 | -14.4715
8.2224 | -34.3568 | -14.4816
8.2285 | -34.3572 | -14.4914
8.2346 | -34.3576 | -14.5014
8.2408 | -34.358 | -14.5115
8.2469 | -34.3584 | -14.5217
8.253 | -34.3588 | -14.5317
8.2591 | -34.3592 | -14.5419
8.2652 | -34.3596 | -14.5521
8.2713 | -34.36 | -14.5622
8.2774 | -34.3604 | -14.5725
8.2835 | -34.3607 | -14.5829
8.2896 | -34.3611 | -14.5929
8.2957 | -34.3615 | -14.6031
8.3018 | -34.3619 | -14.6132
8.3079 | -34.3623 | -14.6232
8.314 | -34.3627 | -14.6333
8.3201 | -34.3631 | -14.6432
8.3262 | -34.3635 | -14.6534
8.3323 | -34.3639 | -14.6634
8.3384 | -34.3643 | -14.6738
8.3445 | -34.3647 | -14.6838
8.3506 | -34.3651 | -14.694
8.3567 | -34.3655 | -14.7042
8.3628 | -34.3659 | -14.7143
8.3689 | -34.3662 | -14.7243
8.375 | -34.3666 | -14.7345
8.3812 | -34.367 | -14.7445
8.3873 | -34.3674 | -14.7546
8.3934 | -34.3678 | -14.7648
8.3995 | -34.3682 | -14.7748
8.4056 | -34.3686 | -14.7849
8.4117 | -34.3689 | -14.7951
8.4178 | -34.3693 | -14.8053
8.4239 | -34.3697 | -14.8153
8.43 | -34.3701 | -14.8255
8.4361 | -34.3704 | -14.8357
8.4422 | -34.3708 | -14.8458
8.4483 | -34.3712 | -14.8558
8.4544 | -34.3716 | -14.8659
8.4605 | -34.372 | -14.8759
8.4666 | -34.3724 | -14.8863
8.4727 | -34.3727 | -14.8964
8.4788 | -34.3731 | -14.9065
8.4849 | -34.3735 | -14.9165
8.491 | -34.3739 | -14.9266
8.4971 | -34.3743 | -14.9364
8.5032 | -34.3747 | -14.9464
8.5093 | -34.3751 | -14.9563
8.5154 | -34.3755 | -14.9665
8.5215 | -34.3759 | -14.9764
8.5277 | -34.3763 | -14.9865
8.5338 | -34.3767 | -14.9965
8.5399 | -34.377 | -15.0066
8.546 | -34.3775 | -15.0165
8.5521 | -34.3778 | -15.0266
8.5582 | -34.3782 | -15.0366
8.5643 | -34.3786 | -15.0464
8.5704 | -34.379 | -15.0562
8.5765 | -34.3794 | -15.0664
8.5826 | -34.3798 | -15.0761
8.5887 | -34.3802 | -15.0862
8.5948 | -34.3806 | -15.0962
8.6009 | -34.381 | -15.1059
8.607 | -34.3814 | -15.1161
8.6131 | -34.3818 | -15.1262
8.6192 | -34.3822 | -15.1361
8.6253 | -34.3826 | -15.1461
8.6314 | -34.383 | -15.1561
8.6375 | -34.3834 | -15.1661
8.6436 | -34.3838 | -15.1762
8.6497 | -34.3841 | -15.186
8.6558 | -34.3845 | -15.196
8.6619 | -34.3849 | -15.206
8.6681 | -34.3853 | -15.2161
8.6742 | -34.3857 | -15.2262
8.6803 | -34.3861 | -15.2363
8.6864 | -34.3865 | -15.246
8.6925 | -34.3869 | -15.2558
8.6986 | -34.3873 | -15.2658
8.7047 | -34.3876 | -15.276
8.7108 | -34.388 | -15.2859
8.7169 | -34.3884 | -15.2959
8.723 | -34.3889 | -15.3058
8.7291 | -34.3893 | -15.3156
8.7352 | -34.3897 | -15.3257
8.7413 | -34.3901 | -15.3358
8.7474 | -34.3904 | -15.3454
8.7535 | -34.3909 | -15.3555
8.7596 | -34.3913 | -15.3655
8.7657 | -34.3917 | -15.3754
8.7718 | -34.392 | -15.3855
8.7779 | -34.3925 | -15.3954
8.784 | -34.3928 | -15.4054
8.7901 | -34.3933 | -15.4152
8.7962 | -34.3937 | -15.4249
8.8023 | -34.394 | -15.4348
8.8084 | -34.3944 | -15.4447
8.8146 | -34.3949 | -15.4547
8.8207 | -34.3952 | -15.4644
8.8268 | -34.3957 | -15.4741
8.8329 | -34.396 | -15.4839
8.839 | -34.3965 | -15.4939
8.8451 | -34.3968 | -15.5037
8.8512 | -34.3973 | -15.5134
8.8573 | -34.3976 | -15.5233
8.8634 | -34.3981 | -15.5331
8.8695 | -34.3985 | -15.5431
8.8756 | -34.3988 | -15.553
8.8817 | -34.3992 | -15.5629
8.8878 | -34.3996 | -15.5731
8.8939 | -34.4 | -15.5829
8.9 | -34.4005 | -15.5928
8.9061 | -34.4008 | -15.6026
8.9122 | -34.4013 | -15.6124
8.9183 | -34.4016 | -15.6224
8.9244 | -34.4021 | -15.6323
8.9305 | -34.4025 | -15.642
8.9366 | -34.4029 | -15.6518
8.9427 | -34.4033 | -15.6618
8.9488 | -34.4037 | -15.6716
8.955 | -34.4041 | -15.6814
8.9611 | -34.4045 | -15.6914
8.9672 | -34.405 | -15.7013
8.9733 | -34.4054 | -15.7109
8.9794 | -34.4058 | -15.7206
8.9855 | -34.4062 | -15.7306
8.9916 | -34.4066 | -15.7406
8.9977 | -34.407 | -15.7504
9.0038 | -34.4075 | -15.7601
9.0099 | -34.4079 | -15.77
9.016 | -34.4083 | -15.7799
9.0221 | -34.4087 | -15.7896
9.0282 | -34.4092 | -15.7993
9.0343 | -34.4096 | -15.8091
9.0404 | -34.41 | -15.8189
9.0465 | -34.4104 | -15.829
9.0526 | -34.4109 | -15.8387
9.0587 | -34.4113 | -15.8486
9.0648 | -34.4117 | -15.8583
9.0709 | -34.4121 | -15.8681
9.077 | -34.4126 | -15.878
9.0831 | -34.413 | -15.888
9.0892 | -34.4134 | -15.8979
9.0953 | -34.4139 | -15.9078
9.1015 | -34.4143 | -15.9176
9.1076 | -34.4147 | -15.9273
9.1137 | -34.4151 | -15.9371
9.1198 | -34.4156 | -15.9469
9.1259 | -34.416 | -15.9567
9.132 | -34.4164 | -15.9667
9.1381 | -34.4169 | -15.9764
9.1442 | -34.4173 | -15.9858
9.1503 | -34.4177 | -15.9957
9.1564 | -34.4182 | -16.0055
9.1625 | -34.4186 | -16.0152
9.1686 | -34.419 | -16.0249
9.1747 | -34.4194 | -16.0348
9.1808 | -34.4199 | -16.0445
9.1869 | -34.4203 | -16.0543
9.193 | -34.4207 | -16.0642
9.1991 | -34.4211 | -16.074
9.2052 | -34.4216 | -16.0839
9.2113 | -34.422 | -16.0936
9.2174 | -34.4225 | -16.1033
9.2235 | -34.423 | -16.1131
9.2296 | -34.4234 | -16.1227
9.2357 | -34.4238 | -16.1326
9.2419 | -34.4243 | -16.1424
9.248 | -34.4248 | -16.1521
9.2541 | -34.4252 | -16.162
9.2602 | -34.4256 | -16.1719
9.2663 | -34.4261 | -16.1814
9.2724 | -34.4265 | -16.191
9.2785 | -34.4269 | -16.2008
9.2846 | -34.4274 | -16.2103
9.2907 | -34.4278 | -16.22
9.2968 | -34.4282 | -16.2299
9.3029 | -34.4287 | -16.2398
9.309 | -34.4291 | -16.2496
9.3151 | -34.4295 | -16.2592
9.3212 | -34.43 | -16.2687
9.3273 | -34.4305 | -16.2784
9.3334 | -34.4309 | -16.2879
9.3395 | -34.4313 | -16.2976
9.3456 | -34.4317 | -16.3073
9.3517 | -34.4321 | -16.3171
9.3578 | -34.4326 | -16.3269
9.3639 | -34.4331 | -16.3365
9.37 | -34.4335 | -16.3463
9.3761 | -34.4339 | -16.3561
9.3822 | -34.4344 | -16.3659
9.3884 | -34.4348 | -16.3753
9.3945 | -34.4353 | -16.3852
9.4006 | -34.4358 | -16.3949
9.4067 | -34.4362 | -16.4048
9.4128 | -34.4367 | -16.4146
9.4189 | -34.4371 | -16.4245
9.425 | -34.4376 | -16.4341
9.4311 | -34.438 | -16.4439
9.4372 | -34.4385 | -16.4536
9.4433 | -34.4389 | -16.4634
9.4494 | -34.4394 | -16.4729
9.4555 | -34.4398 | -16.4827
9.4616 | -34.4403 | -16.4923
9.4677 | -34.4407 | -16.502
9.4738 | -34.4412 | -16.5117
9.4799 | -34.4416 | -16.5215
9.486 | -34.4421 | -16.5312
9.4921 | -34.4425 | -16.5409
9.4982 | -34.443 | -16.5504
9.5043 | -34.4434 | -16.56
9.5104 | -34.4439 | -16.5696
9.5165 | -34.4443 | -16.5793
9.5226 | -34.4448 | -16.5891
9.5288 | -34.4452 | -16.5987
9.5349 | -34.4457 | -16.6084
9.541 | -34.4461 | -16.618
9.5471 | -34.4466 | -16.6276
9.5532 | -34.447 | -16.6374
9.5593 | -34.4475 | -16.6472
9.5654 | -34.4479 | -16.6568
9.5715 | -34.4484 | -16.6665
9.5776 | -34.4488 | -16.6763
9.5837 | -34.4492 | -16.686
9.5898 | -34.4497 | -16.6958
9.5959 | -34.4501 | -16.7055
9.602 | -34.4506 | -16.715
9.6081 | -34.4511 | -16.7248
9.6142 | -34.4516 | -16.7343
9.6203 | -34.452 | -16.744
9.6264 | -34.4524 | -16.7536
9.6325 | -34.4529 | -16.7633
9.6386 | -34.4533 | -16.7729
9.6447 | -34.4538 | -16.7825
9.6508 | -34.4542 | -16.7921
9.6569 | -34.4547 | -16.8017
9.663 | -34.4551 | -16.8114
9.6691 | -34.4556 | -16.8211
9.6753 | -34.4561 | -16.8307
9.6814 | -34.4565 | -16.8402
9.6875 | -34.457 | -16.8496
9.6936 | -34.4574 | -16.8591
9.6997 | -34.4579 | -16.8689
9.7058 | -34.4584 | -16.8784
9.7119 | -34.4588 | -16.8882
9.718 | -34.4593 | -16.8978
9.7241 | -34.4597 | -16.9074
9.7302 | -34.4602 | -16.917
9.7363 | -34.4606 | -16.9265
9.7424 | -34.4611 | -16.936
9.7485 | -34.4615 | -16.9457
9.7546 | -34.462 | -16.9552
9.7607 | -34.4624 | -16.965
9.7668 | -34.4629 | -16.9747
9.7729 | -34.4634 | -16.9842
9.779 | -34.4638 | -16.9937
9.7851 | -34.4643 | -17.0034
9.7912 | -34.4648 | -17.013
9.7973 | -34.4653 | -17.0228
9.8034 | -34.4657 | -17.0322
9.8095 | -34.4662 | -17.042
9.8157 | -34.4666 | -17.0514
9.8218 | -34.4671 | -17.0612
9.8279 | -34.4675 | -17.0707
9.834 | -34.4679 | -17.0804
9.8401 | -34.4684 | -17.0901
9.8462 | -34.4689 | -17.0995
9.8523 | -34.4693 | -17.109
9.8584 | -34.4697 | -17.1186
9.8645 | -34.4702 | -17.1283
9.8706 | -34.4706 | -17.1379
9.8767 | -34.471 | -17.1474
9.8828 | -34.4715 | -17.157
9.8889 | -34.4719 | -17.1663
9.895 | -34.4724 | -17.1759
9.9011 | -34.4728 | -17.1854
9.9072 | -34.4733 | -17.1949
9.9133 | -34.4737 | -17.2046
9.9194 | -34.4742 | -17.2141
9.9255 | -34.4746 | -17.2236
9.9316 | -34.4751 | -17.2332
9.9377 | -34.4755 | -17.2427
9.9438 | -34.4759 | -17.2523
9.9499 | -34.4764 | -17.2622
9.956 | -34.4768 | -17.2717
9.9622 | -34.4773 | -17.2813
9.9683 | -34.4777 | -17.2908
9.9744 | -34.4782 | -17.3006
9.9805 | -34.4787 | -17.3101
9.9866 | -34.4791 | -17.3196
9.9927 | -34.4795 | -17.3291
9.9988 | -34.4799 | -17.3387
10.0049 | -34.4804 | -17.3483
10.011 | -34.4809 | -17.3579
10.0171 | -34.4813 | -17.3674
10.0232 | -34.4818 | -17.3769
10.0293 | -34.4823 | -17.3866
10.0354 | -34.4828 | -17.3961
10.0415 | -34.4832 | -17.4057
10.0476 | -34.4837 | -17.4154
10.0537 | -34.4841 | -17.4251
10.0598 | -34.4846 | -17.4347
10.0659 | -34.485 | -17.4443
10.072 | -34.4854 | -17.4536
10.0781 | -34.4859 | -17.4631
10.0842 | -34.4863 | -17.4727
10.0903 | -34.4868 | -17.4822
10.0964 | -34.4873 | -17.4915
10.1026 | -34.4877 | -17.5011
10.1087 | -34.4881 | -17.5106
10.1148 | -34.4885 | -17.5202
10.1209 | -34.489 | -17.5297
10.127 | -34.4894 | -17.539
10.1331 | -34.4899 | -17.5485
10.1392 | -34.4903 | -17.5577
10.1453 | -34.4908 | -17.5673
10.1514 | -34.4912 | -17.5768
10.1575 | -34.4916 | -17.5864
10.1636 | -34.4921 | -17.5957
10.1697 | -34.4925 | -17.6053
10.1758 | -34.493 | -17.6147
10.1819 | -34.4935 | -17.6242
10.188 | -34.4939 | -17.6337
10.1941 | -34.4943 | -17.6432
10.2002 | -34.4947 | -17.6527
10.2063 | -34.4952 | -17.6622
10.2124 | -34.4956 | -17.6719
10.2185 | -34.4961 | -17.6814
10.2246 | -34.4965 | -17.6908
10.2307 | -34.497 | -17.7004
10.2368 | -34.4974 | -17.7095
10.2429 | -34.4978 | -17.719
10.2491 | -34.4983 | -17.7286
10.2552 | -34.4987 | -17.7381
10.2613 | -34.4992 | -17.7476
10.2674 | -34.4996 | -17.7572
10.2735 | -34.5 | -17.7668
10.2796 | -34.5005 | -17.7762
10.2857 | -34.5009 | -17.7857
10.2918 | -34.5014 | -17.7953
10.2979 | -34.5018 | -17.8047
10.304 | -34.5023 | -17.814
10.3101 | -34.5027 | -17.8235
10.3162 | -34.5032 | -17.8331
10.3223 | -34.5036 | -17.8424
10.3284 | -34.504 | -17.8519
10.3345 | -34.5044 | -17.8613
10.3406 | -34.5049 | -17.8706
10.3467 | -34.5053 | -17.8801
10.3528 | -34.5057 | -17.8896
10.3589 | -34.5062 | -17.8991
10.365 | -34.5066 | -17.9088
10.3711 | -34.5071 | -17.9183
10.3772 | -34.5075 | -17.9278
10.3833 | -34.5079 | -17.9373
10.3895 | -34.5084 | -17.9467
10.3956 | -34.5088 | -17.9562
10.4017 | -34.5093 | -17.9655
10.4078 | -34.5097 | -17.975
10.4139 | -34.5101 | -17.9846
10.42 | -34.5106 | -17.9942
10.4261 | -34.511 | -18.0037
10.4322 | -34.5114 | -18.0132
10.4383 | -34.5119 | -18.0227
10.4444 | -34.5123 | -18.032
10.4505 | -34.5127 | -18.0415
10.4566 | -34.5132 | -18.0508
10.4627 | -34.5136 | -18.0602
10.4688 | -34.514 | -18.0695
10.4749 | -34.5145 | -18.079
10.481 | -34.5149 | -18.0886
10.4871 | -34.5154 | -18.098
10.4932 | -34.5158 | -18.1074
10.4993 | -34.5163 | -18.1169
10.5054 | -34.5167 | -18.1264
10.5115 | -34.5171 | -18.1359
10.5176 | -34.5176 | -18.1452
10.5237 | -34.518 | -18.1549
10.5298 | -34.5185 | -18.164
10.536 | -34.5189 | -18.1735
10.5421 | -34.5193 | -18.1829
10.5482 | -34.5197 | -18.1924
10.5543 | -34.5202 | -18.2018
10.5604 | -34.5206 | -18.2111
10.5665 | -34.521 | -18.2205
10.5726 | -34.5215 | -18.2299
10.5787 | -34.5219 | -18.2393
10.5848 | -34.5223 | -18.2484
10.5909 | -34.5228 | -18.258
10.597 | -34.5232 | -18.2674
10.6031 | -34.5237 | -18.2771
10.6092 | -34.5241 | -18.2865
10.6153 | -34.5245 | -18.2961
10.6214 | -34.525 | -18.3054
10.6275 | -34.5254 | -18.3149
10.6336 | -34.5259 | -18.3243
10.6397 | -34.5263 | -18.3338
10.6458 | -34.5267 | -18.3431
10.6519 | -34.5272 | -18.3526
10.658 | -34.5276 | -18.362
10.6641 | -34.528 | -18.3715
10.6702 | -34.5285 | -18.3809
10.6764 | -34.5289 | -18.3904
10.6825 | -34.5294 | -18.3996
10.6886 | -34.5298 | -18.4089
10.6947 | -34.5302 | -18.4183
10.7008 | -34.5307 | -18.4277
10.7069 | -34.5311 | -18.4371
10.713 | -34.5316 | -18.4465
10.7191 | -34.532 | -18.4559
10.7252 | -34.5325 | -18.4651
10.7313 | -34.5329 | -18.4745
10.7374 | -34.5333 | -18.484
10.7435 | -34.5337 | -18.4936
10.7496 | -34.5342 | -18.5028
10.7557 | -34.5346 | -18.5122
10.7618 | -34.5351 | -18.5214
10.7679 | -34.5355 | -18.5308
10.774 | -34.5359 | -18.5402
10.7801 | -34.5364 | -18.5497
10.7862 | -34.5368 | -18.5589
10.7923 | -34.5372 | -18.5685
10.7984 | -34.5377 | -18.5779
10.8045 | -34.5381 | -18.587
10.8106 | -34.5385 | -18.5966
10.8168 | -34.539 | -18.606
10.8229 | -34.5394 | -18.6154
10.829 | -34.5399 | -18.6248
10.8351 | -34.5403 | -18.634
10.8412 | -34.5408 | -18.6432
10.8473 | -34.5412 | -18.6523
10.8534 | -34.5417 | -18.6618
10.8595 | -34.5421 | -18.6712
10.8656 | -34.5426 | -18.6803
10.8717 | -34.543 | -18.6897
10.8778 | -34.5435 | -18.6991
10.8839 | -34.5439 | -18.7085
10.89 | -34.5444 | -18.718
10.8961 | -34.5448 | -18.7274
10.9022 | -34.5453 | -18.7367
10.9083 | -34.5457 | -18.746
10.9144 | -34.5461 | -18.7552
10.9205 | -34.5465 | -18.7646
10.9266 | -34.5469 | -18.7739
10.9327 | -34.5474 | -18.7832
10.9388 | -34.5478 | -18.7924
10.9449 | -34.5482 | -18.8016
10.951 | -34.5487 | -18.8109
10.9571 | -34.5491 | -18.8203
10.9633 | -34.5495 | -18.8298
10.9694 | -34.5499 | -18.8394
10.9755 | -34.5504 | -18.8488
10.9816 | -34.5508 | -18.8582
10.9877 | -34.5513 | -18.8675
10.9938 | -34.5517 | -18.8769
10.9999 | -34.5522 | -18.8861
11.006 | -34.5526 | -18.8954
11.0121 | -34.5531 | -18.9047
11.0182 | -34.5535 | -18.9141
11.0243 | -34.5539 | -18.9234
11.0304 | -34.5544 | -18.9327
11.0365 | -34.5548 | -18.9419
11.0426 | -34.5553 | -18.9512
11.0487 | -34.5557 | -18.9607
11.0548 | -34.5562 | -18.9699
11.0609 | -34.5567 | -18.9793
11.067 | -34.5571 | -18.9887
11.0731 | -34.5576 | -18.9979
11.0792 | -34.558 | -19.0073
11.0853 | -34.5585 | -19.0166
11.0914 | -34.5589 | -19.0261
11.0975 | -34.5593 | -19.0354
11.1037 | -34.5598 | -19.0447
11.1098 | -34.5602 | -19.0541
11.1159 | -34.5607 | -19.0632
11.122 | -34.5611 | -19.0724
11.1281 | -34.5615 | -19.0817
11.1342 | -34.562 | -19.091
11.1403 | -34.5624 | -19.1004
11.1464 | -34.5629 | -19.1098
11.1525 | -34.5633 | -19.119
11.1586 | -34.5637 | -19.1285
11.1647 | -34.5642 | -19.1378
11.1708 | -34.5646 | -19.1471
11.1769 | -34.5651 | -19.1565
11.183 | -34.5655 | -19.1656
11.1891 | -34.566 | -19.175
11.1952 | -34.5665 | -19.1844
11.2013 | -34.5669 | -19.1936
11.2074 | -34.5674 | -19.2029
11.2135 | -34.5678 | -19.2124
11.2196 | -34.5683 | -19.2217
11.2257 | -34.5687 | -19.2311
11.2318 | -34.5692 | -19.2403
11.2379 | -34.5697 | -19.2499
11.244 | -34.5702 | -19.2593
11.2502 | -34.5706 | -19.2687
11.2563 | -34.5711 | -19.2781
11.2624 | -34.5716 | -19.2874
11.2685 | -34.572 | -19.2966
11.2746 | -34.5724 | -19.3059
11.2807 | -34.5729 | -19.3153
11.2868 | -34.5734 | -19.3248
11.2929 | -34.5739 | -19.3337
11.299 | -34.5743 | -19.3432
11.3051 | -34.5748 | -19.3527
11.3112 | -34.5752 | -19.362
11.3173 | -34.5757 | -19.3712
11.3234 | -34.5761 | -19.3804
11.3295 | -34.5766 | -19.3898
11.3356 | -34.577 | -19.3991
11.3417 | -34.5775 | -19.4082
11.3478 | -34.5779 | -19.4177
11.3539 | -34.5783 | -19.4271
11.36 | -34.5788 | -19.4362
11.3661 | -34.5793 | -19.4453
11.3722 | -34.5797 | -19.4544
11.3783 | -34.5802 | -19.4638
11.3844 | -34.5807 | -19.4732
11.3906 | -34.5812 | -19.4824
11.3967 | -34.5816 | -19.4917
11.4028 | -34.5821 | -19.5012
11.4089 | -34.5825 | -19.5104
11.415 | -34.583 | -19.5195
11.4211 | -34.5834 | -19.5287
11.4272 | -34.5839 | -19.5381
11.4333 | -34.5843 | -19.5472
11.4394 | -34.5848 | -19.5565
11.4455 | -34.5852 | -19.5657
11.4516 | -34.5857 | -19.575
11.4577 | -34.5861 | -19.5842
11.4638 | -34.5866 | -19.5933
11.4699 | -34.5871 | -19.6026
11.476 | -34.5875 | -19.6121
11.4821 | -34.588 | -19.6213
11.4882 | -34.5885 | -19.6307
11.4943 | -34.5889 | -19.6401
11.5004 | -34.5894 | -19.6491
11.5065 | -34.5899 | -19.6583
11.5126 | -34.5904 | -19.6675
11.5187 | -34.5908 | -19.6769
11.5248 | -34.5912 | -19.6863
11.5309 | -34.5917 | -19.6956
11.5371 | -34.5922 | -19.7047
11.5432 | -34.5927 | -19.7141
11.5493 | -34.5931 | -19.7235
11.5554 | -34.5936 | -19.7329
11.5615 | -34.5941 | -19.7422
11.5676 | -34.5945 | -19.7515
11.5737 | -34.595 | -19.7607
11.5798 | -34.5955 | -19.7698
11.5859 | -34.596 | -19.7791
11.592 | -34.5964 | -19.7884
11.5981 | -34.5969 | -19.7977
11.6042 | -34.5974 | -19.807
11.6103 | -34.5978 | -19.8161
11.6164 | -34.5983 | -19.8253
11.6225 | -34.5987 | -19.8345
11.6286 | -34.5992 | -19.8437
11.6347 | -34.5996 | -19.8531
11.6408 | -34.6001 | -19.8623
11.6469 | -34.6006 | -19.8716
11.653 | -34.6011 | -19.8808
11.6591 | -34.6015 | -19.8901
11.6652 | -34.602 | -19.8992
11.6713 | -34.6025 | -19.9084
11.6775 | -34.603 | -19.9176
11.6836 | -34.6034 | -19.9269
11.6897 | -34.6039 | -19.9362
11.6958 | -34.6044 | -19.9456
11.7019 | -34.6048 | -19.9549
11.708 | -34.6053 | -19.9642
11.7141 | -34.6058 | -19.9733
11.7202 | -34.6063 | -19.9825
11.7263 | -34.6067 | -19.9918
11.7324 | -34.6072 | -20.0013
11.7385 | -34.6076 | -20.0106
11.7446 | -34.6081 | -20.0198
11.7507 | -34.6086 | -20.0289
11.7568 | -34.609 | -20.0381
11.7629 | -34.6095 | -20.0475
11.769 | -34.61 | -20.0565
11.7751 | -34.6104 | -20.0657
11.7812 | -34.6109 | -20.0749
11.7873 | -34.6114 | -20.0842
11.7934 | -34.6119 | -20.0934
11.7995 | -34.6124 | -20.1026
11.8056 | -34.6128 | -20.1118
11.8117 | -34.6133 | -20.121
11.8178 | -34.6138 | -20.1301
11.824 | -34.6143 | -20.1392
11.8301 | -34.6148 | -20.1483
11.8362 | -34.6152 | -20.1574
11.8423 | -34.6157 | -20.1665
11.8484 | -34.6162 | -20.1758
11.8545 | -34.6167 | -20.1848
11.8606 | -34.6172 | -20.1938
11.8667 | -34.6176 | -20.2031
11.8728 | -34.6181 | -20.2122
11.8789 | -34.6185 | -20.2213
11.885 | -34.619 | -20.2306
11.8911 | -34.6195 | -20.2398
11.8972 | -34.62 | -20.2488
11.9033 | -34.6205 | -20.2579
11.9094 | -34.6209 | -20.2669
11.9155 | -34.6214 | -20.2761
11.9216 | -34.6219 | -20.2851
11.9277 | -34.6224 | -20.2941
11.9338 | -34.6229 | -20.3033
11.9399 | -34.6234 | -20.3125
11.946 | -34.6239 | -20.3216
11.9521 | -34.6243 | -20.3306
11.9582 | -34.6248 | -20.3396
11.9644 | -34.6253 | -20.3486
11.9705 | -34.6258 | -20.3577
11.9766 | -34.6263 | -20.3668
11.9827 | -34.6268 | -20.3759
11.9888 | -34.6273 | -20.385
11.9949 | -34.6278 | -20.3938
12.001 | -34.6283 | -20.403
12.0071 | -34.6288 | -20.4119
12.0132 | -34.6292 | -20.4209
12.0193 | -34.6297 | -20.4301
12.0254 | -34.6302 | -20.4391
12.0315 | -34.6307 | -20.448
12.0376 | -34.6311 | -20.457
12.0437 | -34.6316 | -20.466
12.0498 | -34.6321 | -20.4752
12.0559 | -34.6325 | -20.4842
12.062 | -34.6331 | -20.4931
12.0681 | -34.6336 | -20.5023
12.0742 | -34.634 | -20.5112
12.0803 | -34.6345 | -20.5205
12.0864 | -34.635 | -20.5296
12.0925 | -34.6355 | -20.5387
12.0986 | -34.636 | -20.5477
12.1047 | -34.6365 | -20.5568
12.1109 | -34.637 | -20.5657
12.117 | -34.6374 | -20.5747
12.1231 | -34.6379 | -20.5838
12.1292 | -34.6384 | -20.5927
12.1353 | -34.6389 | -20.6017
12.1414 | -34.6394 | -20.6108
12.1475 | -34.6399 | -20.62
12.1536 | -34.6404 | -20.6292
12.1597 | -34.6408 | -20.6382
12.1658 | -34.6414 | -20.6471
12.1719 | -34.6418 | -20.656
12.178 | -34.6423 | -20.6651
12.1841 | -34.6428 | -20.6743
12.1902 | -34.6433 | -20.6834
12.1963 | -34.6438 | -20.6924
12.2024 | -34.6443 | -20.7013
12.2085 | -34.6448 | -20.7103
12.2146 | -34.6453 | -20.7194
12.2207 | -34.6458 | -20.7284
12.2268 | -34.6463 | -20.7376
12.2329 | -34.6468 | -20.7465
12.239 | -34.6473 | -20.7552
12.2451 | -34.6478 | -20.7642
12.2513 | -34.6482 | -20.7733
12.2574 | -34.6487 | -20.7823
12.2635 | -34.6492 | -20.7913
12.2696 | -34.6497 | -20.8002
12.2757 | -34.6502 | -20.809
12.2818 | -34.6507 | -20.8181
12.2879 | -34.6512 | -20.827
12.294 | -34.6517 | -20.8359
12.3001 | -34.6522 | -20.8449
12.3062 | -34.6527 | -20.854
12.3123 | -34.6532 | -20.8626
12.3184 | -34.6537 | -20.8715
12.3245 | -34.6542 | -20.8804
12.3306 | -34.6547 | -20.8893
12.3367 | -34.6552 | -20.8982
12.3428 | -34.6557 | -20.9072
12.3489 | -34.6562 | -20.9158
12.355 | -34.6567 | -20.9249
12.3611 | -34.6572 | -20.934
12.3672 | -34.6576 | -20.9429
12.3733 | -34.6582 | -20.9519
12.3794 | -34.6587 | -20.9605
12.3855 | -34.6591 | -20.9694
12.3916 | -34.6596 | -20.9784
12.3978 | -34.6602 | -20.9875
12.4039 | -34.6606 | -20.9965
12.41 | -34.6611 | -21.0055
12.4161 | -34.6617 | -21.014
12.4222 | -34.6622 | -21.023
12.4283 | -34.6626 | -21.0318
12.4344 | -34.6631 | -21.0409
12.4405 | -34.6636 | -21.0499
12.4466 | -34.6641 | -21.0588
12.4527 | -34.6646 | -21.0676
12.4588 | -34.6651 | -21.0765
12.4649 | -34.6656 | -21.0855
12.471 | -34.6661 | -21.0946
12.4771 | -34.6666 | -21.1035
12.4832 | -34.6671 | -21.1124
12.4893 | -34.6676 | -21.1211
12.4954 | -34.6681 | -21.1299
12.5015 | -34.6686 | -21.1388
12.5076 | -34.6691 | -21.1478
12.5137 | -34.6696 | -21.1567
12.5198 | -34.6701 | -21.1654
12.5259 | -34.6706 | -21.1744
12.532 | -34.6711 | -21.1832
12.5382 | -34.6716 | -21.1923
12.5443 | -34.6721 | -21.2012
12.5504 | -34.6726 | -21.21
12.5565 | -34.6731 | -21.2188
12.5626 | -34.6737 | -21.2278
12.5687 | -34.6741 | -21.2367
12.5748 | -34.6746 | -21.2456
12.5809 | -34.6751 | -21.2545
12.587 | -34.6756 | -21.2635
12.5931 | -34.6761 | -21.2721
12.5992 | -34.6767 | -21.2809
12.6053 | -34.6771 | -21.2897
12.6114 | -34.6776 | -21.2985
12.6175 | -34.6782 | -21.3074
12.6236 | -34.6787 | -21.3163
12.6297 | -34.6792 | -21.325
12.6358 | -34.6797 | -21.3336
12.6419 | -34.6802 | -21.3426
12.648 | -34.6807 | -21.3515
12.6541 | -34.6811 | -21.3602
12.6602 | -34.6817 | -21.3687
12.6663 | -34.6822 | -21.3776
12.6724 | -34.6827 | -21.3863
12.6785 | -34.6832 | -21.3952
12.6847 | -34.6837 | -21.404
12.6908 | -34.6842 | -21.4128
12.6969 | -34.6848 | -21.4215
12.703 | -34.6852 | -21.4304
12.7091 | -34.6857 | -21.4392
12.7152 | -34.6862 | -21.4479
12.7213 | -34.6867 | -21.4568
12.7274 | -34.6872 | -21.4657
12.7335 | -34.6878 | -21.4742
12.7396 | -34.6883 | -21.4829
12.7457 | -34.6888 | -21.4918
12.7518 | -34.6892 | -21.5004
12.7579 | -34.6897 | -21.5094
12.764 | -34.6902 | -21.5183
12.7701 | -34.6907 | -21.527
12.7762 | -34.6913 | -21.5356
12.7823 | -34.6918 | -21.5445
12.7884 | -34.6923 | -21.5533
12.7945 | -34.6928 | -21.5624
12.8006 | -34.6933 | -21.571
12.8067 | -34.6938 | -21.5797
12.8128 | -34.6943 | -21.5883
12.8189 | -34.6948 | -21.5971
12.8251 | -34.6953 | -21.606
12.8312 | -34.6958 | -21.6149
12.8373 | -34.6963 | -21.6235
12.8434 | -34.6969 | -21.6323
12.8495 | -34.6974 | -21.641
12.8556 | -34.6978 | -21.6497
12.8617 | -34.6984 | -21.6585
12.8678 | -34.6989 | -21.6674
12.8739 | -34.6994 | -21.676
12.88 | -34.6999 | -21.6847
12.8861 | -34.7004 | -21.6934
12.8922 | -34.7009 | -21.7023
12.8983 | -34.7014 | -21.7112
12.9044 | -34.7019 | -21.7197
12.9105 | -34.7024 | -21.7283
12.9166 | -34.7029 | -21.737
12.9227 | -34.7034 | -21.7457
12.9288 | -34.7039 | -21.7545
12.9349 | -34.7044 | -21.7633
12.941 | -34.7049 | -21.772
12.9471 | -34.7054 | -21.7807
12.9532 | -34.706 | -21.7894
12.9593 | -34.7065 | -21.7983
12.9654 | -34.707 | -21.8072
12.9716 | -34.7075 | -21.816
12.9777 | -34.708 | -21.8248
12.9838 | -34.7085 | -21.8334
12.9899 | -34.709 | -21.8421
12.996 | -34.7095 | -21.8509
13.0021 | -34.7101 | -21.8596
13.0082 | -34.7106 | -21.8681
13.0143 | -34.7112 | -21.8764
13.0204 | -34.7116 | -21.8851
13.0265 | -34.7121 | -21.8937
13.0326 | -34.7126 | -21.9024
13.0387 | -34.7132 | -21.9111
13.0448 | -34.7137 | -21.9198
13.0509 | -34.7142 | -21.9285
13.057 | -34.7147 | -21.9371
13.0631 | -34.7152 | -21.9457
13.0692 | -34.7157 | -21.9542
13.0753 | -34.7163 | -21.9631
13.0814 | -34.7167 | -21.9718
13.0875 | -34.7173 | -21.9803
13.0936 | -34.7177 | -21.9891
13.0997 | -34.7183 | -21.9977
13.1058 | -34.7188 | -22.0062
13.112 | -34.7193 | -22.0148
13.1181 | -34.7198 | -22.0233
13.1242 | -34.7203 | -22.0322
13.1303 | -34.7208 | -22.0408
13.1364 | -34.7213 | -22.0494
13.1425 | -34.7218 | -22.058
13.1486 | -34.7223 | -22.0667
13.1547 | -34.7228 | -22.0752
13.1608 | -34.7233 | -22.0839
13.1669 | -34.7239 | -22.0924
13.173 | -34.7244 | -22.1012
13.1791 | -34.7249 | -22.1097
13.1852 | -34.7254 | -22.1183
13.1913 | -34.7259 | -22.1268
13.1974 | -34.7264 | -22.1355
13.2035 | -34.7269 | -22.1442
13.2096 | -34.7275 | -22.1528
13.2157 | -34.728 | -22.1613
13.2218 | -34.7285 | -22.1699
13.2279 | -34.729 | -22.1782
13.234 | -34.7295 | -22.1869
13.2401 | -34.73 | -22.1955
13.2462 | -34.7305 | -22.2041
13.2524 | -34.7311 | -22.2129
13.2585 | -34.7316 | -22.2214
13.2646 | -34.7321 | -22.23
13.2707 | -34.7326 | -22.2387
13.2768 | -34.7331 | -22.2475
13.2829 | -34.7336 | -22.256
13.289 | -34.7341 | -22.2644
13.2951 | -34.7347 | -22.2729
13.3012 | -34.7352 | -22.2816
13.3073 | -34.7357 | -22.2902
13.3134 | -34.7362 | -22.2988
13.3195 | -34.7367 | -22.3072
13.3256 | -34.7372 | -22.3158
13.3317 | -34.7377 | -22.3243
13.3378 | -34.7382 | -22.333
13.3439 | -34.7387 | -22.3416
13.35 | -34.7393 | -22.3503
13.3561 | -34.7398 | -22.3587
13.3622 | -34.7403 | -22.3671
13.3683 | -34.7408 | -22.3756
13.3744 | -34.7413 | -22.384
13.3805 | -34.7418 | -22.3928
13.3866 | -34.7423 | -22.4013
13.3927 | -34.7428 | -22.4098
13.3989 | -34.7433 | -22.4183
13.405 | -34.7438 | -22.4269
13.4111 | -34.7443 | -22.4353
13.4172 | -34.7448 | -22.4439
13.4233 | -34.7454 | -22.4526
13.4294 | -34.7459 | -22.4614
13.4355 | -34.7464 | -22.4697
13.4416 | -34.7469 | -22.4782
13.4477 | -34.7474 | -22.4867
13.4538 | -34.7479 | -22.4952
13.4599 | -34.7484 | -22.5039
13.466 | -34.7489 | -22.5125
13.4721 | -34.7494 | -22.5208
13.4782 | -34.7499 | -22.5293
13.4843 | -34.7505 | -22.5379
13.4904 | -34.751 | -22.5463
13.4965 | -34.7515 | -22.5549
13.5026 | -34.752 | -22.5636
13.5087 | -34.7525 | -22.572
13.5148 | -34.753 | -22.5807
13.5209 | -34.7535 | -22.5894
13.527 | -34.754 | -22.5979
13.5331 | -34.7545 | -22.6065
13.5393 | -34.7551 | -22.6149
13.5454 | -34.7556 | -22.6233
13.5515 | -34.7561 | -22.6321
13.5576 | -34.7566 | -22.6406
13.5637 | -34.7571 | -22.6493
13.5698 | -34.7576 | -22.6579
13.5759 | -34.7581 | -22.666
13.582 | -34.7586 | -22.6744
13.5881 | -34.7592 | -22.6829
13.5942 | -34.7597 | -22.6916
13.6003 | -34.7602 | -22.7001
13.6064 | -34.7607 | -22.7088
13.6125 | -34.7612 | -22.7173
13.6186 | -34.7618 | -22.7258
13.6247 | -34.7623 | -22.7343
13.6308 | -34.7628 | -22.7429
13.6369 | -34.7633 | -22.7515
13.643 | -34.7638 | -22.7601
13.6491 | -34.7643 | -22.7685
13.6552 | -34.7649 | -22.777
13.6613 | -34.7654 | -22.7855
13.6674 | -34.7658 | -22.794
13.6735 | -34.7664 | -22.8025
13.6796 | -34.7669 | -22.811
13.6858 | -34.7674 | -22.8195
13.6919 | -34.7679 | -22.828
13.698 | -34.7684 | -22.8366
13.7041 | -34.769 | -22.845
13.7102 | -34.7695 | -22.8537
13.7163 | -34.77 | -22.8622
13.7224 | -34.7705 | -22.8708
13.7285 | -34.7711 | -22.8793
13.7346 | -34.7716 | -22.8878
13.7407 | -34.7721 | -22.8963
13.7468 | -34.7726 | -22.9048
13.7529 | -34.7731 | -22.9132
13.759 | -34.7736 | -22.9218
13.7651 | -34.7741 | -22.9303
13.7712 | -34.7746 | -22.9389
13.7773 | -34.7752 | -22.9475
13.7834 | -34.7757 | -22.956
13.7895 | -34.7762 | -22.9644
13.7956 | -34.7767 | -22.973
13.8017 | -34.7772 | -22.9814
13.8078 | -34.7777 | -22.9899
13.8139 | -34.7782 | -22.9983
13.82 | -34.7787 | -23.0068
13.8262 | -34.7793 | -23.0153
13.8323 | -34.7798 | -23.0239
13.8384 | -34.7803 | -23.0324
13.8445 | -34.7808 | -23.041
13.8506 | -34.7813 | -23.0495
13.8567 | -34.7819 | -23.058
13.8628 | -34.7824 | -23.0664
13.8689 | -34.7829 | -23.0749
13.875 | -34.7834 | -23.0833
13.8811 | -34.7839 | -23.0919
13.8872 | -34.7844 | -23.1003
13.8933 | -34.7849 | -23.1086
13.8994 | -34.7854 | -23.117
13.9055 | -34.786 | -23.1257
13.9116 | -34.7865 | -23.1341
13.9177 | -34.787 | -23.1426
13.9238 | -34.7875 | -23.151
13.9299 | -34.788 | -23.1594
13.936 | -34.7885 | -23.1678
13.9421 | -34.789 | -23.1763
13.9482 | -34.7895 | -23.1847
13.9543 | -34.79 | -23.1933
13.9604 | -34.7905 | -23.2018
13.9665 | -34.7911 | -23.2102
13.9727 | -34.7916 | -23.2189
13.9788 | -34.7921 | -23.2273
13.9849 | -34.7926 | -23.2357
13.991 | -34.7932 | -23.2441
13.9971 | -34.7937 | -23.2524
14.0032 | -34.7942 | -23.2609
14.0093 | -34.7947 | -23.2694
14.0154 | -34.7952 | -23.2778
14.0215 | -34.7957 | -23.2861
14.0276 | -34.7963 | -23.2947
14.0337 | -34.7968 | -23.3031
14.0398 | -34.7973 | -23.3115
14.0459 | -34.7978 | -23.3199
14.052 | -34.7984 | -23.3284
14.0581 | -34.7988 | -23.3371
14.0642 | -34.7993 | -23.3456
14.0703 | -34.7999 | -23.3541
14.0764 | -34.8004 | -23.3626
14.0825 | -34.8009 | -23.3711
14.0886 | -34.8014 | -23.3796
14.0947 | -34.802 | -23.388
14.1008 | -34.8025 | -23.3964
14.1069 | -34.803 | -23.4047
14.1131 | -34.8035 | -23.4133
14.1192 | -34.804 | -23.4216
14.1253 | -34.8045 | -23.4298
14.1314 | -34.8051 | -23.4383
14.1375 | -34.8056 | -23.4465
14.1436 | -34.8061 | -23.4549
14.1497 | -34.8067 | -23.4634
14.1558 | -34.8072 | -23.4718
14.1619 | -34.8077 | -23.4802
14.168 | -34.8082 | -23.4885
14.1741 | -34.8088 | -23.4968
14.1802 | -34.8093 | -23.5052
14.1863 | -34.8098 | -23.5136
14.1924 | -34.8103 | -23.5221
14.1985 | -34.8109 | -23.5305
14.2046 | -34.8114 | -23.539
14.2107 | -34.8119 | -23.5471
14.2168 | -34.8124 | -23.5556
14.2229 | -34.8129 | -23.5639
14.229 | -34.8135 | -23.5724
14.2351 | -34.814 | -23.5808
14.2412 | -34.8145 | -23.5892
14.2473 | -34.815 | -23.5975
14.2534 | -34.8155 | -23.6058
14.2596 | -34.816 | -23.6141
14.2657 | -34.8165 | -23.6225
14.2718 | -34.817 | -23.631
14.2779 | -34.8176 | -23.6395
14.284 | -34.8181 | -23.6479
14.2901 | -34.8186 | -23.6564
14.2962 | -34.8192 | -23.6647
14.3023 | -34.8197 | -23.6732
14.3084 | -34.8202 | -23.6817
14.3145 | -34.8207 | -23.69
14.3206 | -34.8212 | -23.6986
14.3267 | -34.8218 | -23.707
14.3328 | -34.8223 | -23.7154
14.3389 | -34.8228 | -23.7236
14.345 | -34.8233 | -23.732
14.3511 | -34.8238 | -23.7404
14.3572 | -34.8244 | -23.7488
14.3633 | -34.8249 | -23.7573
14.3694 | -34.8254 | -23.7657
14.3755 | -34.826 | -23.774
14.3816 | -34.8265 | -23.7825
14.3877 | -34.827 | -23.7908
14.3938 | -34.8276 | -23.7994
14.4 | -34.8281 | -23.8077
14.4061 | -34.8286 | -23.816
14.4122 | -34.8292 | -23.8245
14.4183 | -34.8297 | -23.8328
14.4244 | -34.8302 | -23.8412
14.4305 | -34.8308 | -23.8496
14.4366 | -34.8313 | -23.8581
14.4427 | -34.8318 | -23.8666
14.4488 | -34.8323 | -23.8751
14.4549 | -34.8329 | -23.8835
14.461 | -34.8334 | -23.8917
14.4671 | -34.8339 | -23.9001
14.4732 | -34.8344 | -23.9086
14.4793 | -34.835 | -23.917
14.4854 | -34.8355 | -23.9253
14.4915 | -34.836 | -23.9335
14.4976 | -34.8366 | -23.9417
14.5037 | -34.8371 | -23.9499
14.5098 | -34.8376 | -23.9583
14.5159 | -34.8381 | -23.9667
14.522 | -34.8387 | -23.975
14.5281 | -34.8392 | -23.9831
14.5342 | -34.8397 | -23.9916
14.5403 | -34.8403 | -23.9999
14.5465 | -34.8408 | -24.0084
14.5526 | -34.8413 | -24.0169
14.5587 | -34.8418 | -24.0253
14.5648 | -34.8424 | -24.0334
14.5709 | -34.8429 | -24.0417
14.577 | -34.8434 | -24.05
14.5831 | -34.8439 | -24.0583
14.5892 | -34.8444 | -24.0667
14.5953 | -34.845 | -24.0748
14.6014 | -34.8455 | -24.0832
14.6075 | -34.846 | -24.0915
14.6136 | -34.8465 | -24.0999
14.6197 | -34.847 | -24.1084
14.6258 | -34.8476 | -24.1168
14.6319 | -34.8481 | -24.125
14.638 | -34.8486 | -24.1333
14.6441 | -34.8491 | -24.1417
14.6502 | -34.8497 | -24.1502
14.6563 | -34.8502 | -24.1587
14.6624 | -34.8507 | -24.1671
14.6685 | -34.8513 | -24.1752
14.6746 | -34.8518 | -24.1836
14.6807 | -34.8524 | -24.1918
14.6869 | -34.8529 | -24.2001
14.693 | -34.8534 | -24.2085
14.6991 | -34.854 | -24.2169
14.7052 | -34.8545 | -24.2249
14.7113 | -34.8551 | -24.2334
14.7174 | -34.8556 | -24.2417
14.7235 | -34.8561 | -24.2498
14.7296 | -34.8566 | -24.258
14.7357 | -34.8572 | -24.2663
14.7418 | -34.8577 | -24.2746
14.7479 | -34.8583 | -24.2828
14.754 | -34.8588 | -24.2912
14.7601 | -34.8593 | -24.2994
14.7662 | -34.8599 | -24.3077
14.7723 | -34.8604 | -24.3158
14.7784 | -34.8609 | -24.3242
14.7845 | -34.8615 | -24.3325
14.7906 | -34.862 | -24.3408
14.7967 | -34.8625 | -24.349
14.8028 | -34.863 | -24.3573
14.8089 | -34.8636 | -24.3656
14.815 | -34.8641 | -24.3739
14.8211 | -34.8646 | -24.3821
14.8272 | -34.8652 | -24.3905
14.8334 | -34.8657 | -24.3987
14.8395 | -34.8663 | -24.4069
14.8456 | -34.8668 | -24.4151
14.8517 | -34.8673 | -24.4236
14.8578 | -34.8678 | -24.432
14.8639 | -34.8684 | -24.4404
14.87 | -34.8689 | -24.4487
14.8761 | -34.8695 | -24.4568
14.8822 | -34.87 | -24.4652
14.8883 | -34.8705 | -24.4735
14.8944 | -34.871 | -24.4816
14.9005 | -34.8716 | -24.4898
14.9066 | -34.8721 | -24.4981
14.9127 | -34.8726 | -24.5063
14.9188 | -34.8732 | -24.5146
14.9249 | -34.8737 | -24.5229
14.931 | -34.8742 | -24.5311
14.9371 | -34.8748 | -24.5394
14.9432 | -34.8753 | -24.5477
14.9493 | -34.8759 | -24.556
14.9554 | -34.8764 | -24.5643
14.9615 | -34.8769 | -24.5726
14.9676 | -34.8775 | -24.5812
14.9738 | -34.878 | -24.5894
14.9799 | -34.8785 | -24.5978
14.986 | -34.8791 | -24.606
14.9921 | -34.8796 | -24.6143
14.9982 | -34.8801 | -24.6225
15.0043 | -34.8806 | -24.6306
15.0104 | -34.8812 | -24.6391
15.0165 | -34.8817 | -24.6475
15.0226 | -34.8823 | -24.6556
15.0287 | -34.8828 | -24.6639
15.0348 | -34.8833 | -24.6722
15.0409 | -34.8839 | -24.6804
15.047 | -34.8844 | -24.6886
15.0531 | -34.885 | -24.6968
15.0592 | -34.8855 | -24.7052
15.0653 | -34.886 | -24.7134
15.0714 | -34.8866 | -24.7216
15.0775 | -34.8871 | -24.7298
15.0836 | -34.8877 | -24.7381
15.0897 | -34.8883 | -24.7461
15.0958 | -34.8888 | -24.7544
15.1019 | -34.8893 | -24.7626
15.108 | -34.8899 | -24.7707
15.1141 | -34.8904 | -24.7789
15.1203 | -34.891 | -24.7873
15.1264 | -34.8915 | -24.7954
15.1325 | -34.892 | -24.8037
15.1386 | -34.8926 | -24.8119
15.1447 | -34.8931 | -24.8202
15.1508 | -34.8937 | -24.8284
15.1569 | -34.8942 | -24.8367
15.163 | -34.8948 | -24.8446
15.1691 | -34.8953 | -24.8529
15.1752 | -34.8959 | -24.8611
15.1813 | -34.8964 | -24.8693
15.1874 | -34.8969 | -24.8775
15.1935 | -34.8975 | -24.8855
15.1996 | -34.898 | -24.8936
15.2057 | -34.8986 | -24.9018
15.2118 | -34.8991 | -24.9101
15.2179 | -34.8996 | -24.9184
15.224 | -34.9002 | -24.9266
15.2301 | -34.9007 | -24.9345
15.2362 | -34.9013 | -24.9428
15.2423 | -34.9018 | -24.9511
15.2484 | -34.9024 | -24.9593
15.2545 | -34.9029 | -24.9675
15.2607 | -34.9035 | -24.9757
15.2668 | -34.904 | -24.9836
15.2729 | -34.9046 | -24.9917
15.279 | -34.9051 | -25.0001
15.2851 | -34.9057 | -25.0083
15.2912 | -34.9062 | -25.0165
15.2973 | -34.9068 | -25.0248
15.3034 | -34.9073 | -25.0328
15.3095 | -34.9079 | -25.041
15.3156 | -34.9084 | -25.0492
15.3217 | -34.909 | -25.0573
15.3278 | -34.9095 | -25.0656
15.3339 | -34.9101 | -25.0735
15.34 | -34.9107 | -25.0815
15.3461 | -34.9112 | -25.0896
15.3522 | -34.9118 | -25.0978
15.3583 | -34.9123 | -25.1061
15.3644 | -34.9128 | -25.1142
15.3705 | -34.9134 | -25.122
15.3766 | -34.9139 | -25.1301
15.3827 | -34.9145 | -25.1382
15.3888 | -34.915 | -25.1463
15.3949 | -34.9156 | -25.1544
15.401 | -34.9161 | -25.1627
15.4072 | -34.9167 | -25.1707
15.4133 | -34.9172 | -25.1789
15.4194 | -34.9177 | -25.1872
15.4255 | -34.9183 | -25.1952
15.4316 | -34.9188 | -25.2035
15.4377 | -34.9194 | -25.2118
15.4438 | -34.9199 | -25.2199
15.4499 | -34.9205 | -25.2279
15.456 | -34.921 | -25.236
15.4621 | -34.9215 | -25.2442
15.4682 | -34.9221 | -25.2523
15.4743 | -34.9227 | -25.2603
15.4804 | -34.9232 | -25.2685
15.4865 | -34.9237 | -25.2765
15.4926 | -34.9243 | -25.2848
15.4987 | -34.9248 | -25.293
15.5048 | -34.9254 | -25.301
15.5109 | -34.926 | -25.309
15.517 | -34.9266 | -25.3172
15.5231 | -34.9271 | -25.3254
15.5292 | -34.9277 | -25.3333
15.5353 | -34.9282 | -25.3416
15.5414 | -34.9288 | -25.3497
15.5476 | -34.9293 | -25.3576
15.5537 | -34.9299 | -25.3658
15.5598 | -34.9305 | -25.3741
15.5659 | -34.931 | -25.3822
15.572 | -34.9316 | -25.3902
15.5781 | -34.9322 | -25.3982
15.5842 | -34.9327 | -25.4061
15.5903 | -34.9332 | -25.4141
15.5964 | -34.9338 | -25.4221
15.6025 | -34.9344 | -25.4302
15.6086 | -34.9349 | -25.4384
15.6147 | -34.9354 | -25.4465
15.6208 | -34.936 | -25.4544
15.6269 | -34.9366 | -25.4626
15.633 | -34.9371 | -25.4707
15.6391 | -34.9377 | -25.4787
15.6452 | -34.9382 | -25.4869
15.6513 | -34.9388 | -25.4947
15.6574 | -34.9393 | -25.5028
15.6635 | -34.9399 | -25.5109
15.6696 | -34.9405 | -25.5191
15.6757 | -34.941 | -25.5271
15.6818 | -34.9415 | -25.5351
15.688 | -34.9421 | -25.543
15.6941 | -34.9426 | -25.5511
15.7002 | -34.9431 | -25.5591
15.7063 | -34.9437 | -25.5671
15.7124 | -34.9442 | -25.5753
15.7185 | -34.9448 | -25.5834
15.7246 | -34.9453 | -25.5913
15.7307 | -34.9459 | -25.5992
15.7368 | -34.9464 | -25.6073
15.7429 | -34.947 | -25.6152
15.749 | -34.9476 | -25.6233
15.7551 | -34.9481 | -25.6314
15.7612 | -34.9487 | -25.6394
15.7673 | -34.9493 | -25.6474
15.7734 | -34.9498 | -25.6555
15.7795 | -34.9503 | -25.6635
15.7856 | -34.9509 | -25.6713
15.7917 | -34.9515 | -25.6792
15.7978 | -34.952 | -25.6874
15.8039 | -34.9526 | -25.6955
15.81 | -34.9531 | -25.7035
15.8161 | -34.9537 | -25.7115
15.8222 | -34.9542 | -25.7196
15.8283 | -34.9548 | -25.7273
15.8345 | -34.9553 | -25.7353
15.8406 | -34.9559 | -25.7432
15.8467 | -34.9564 | -25.7513
15.8528 | -34.957 | -25.7593
15.8589 | -34.9576 | -25.7672
15.865 | -34.9582 | -25.7749
15.8711 | -34.9587 | -25.7829
15.8772 | -34.9593 | -25.7908
15.8833 | -34.9598 | -25.7989
15.8894 | -34.9604 | -25.807
15.8955 | -34.9609 | -25.815
15.9016 | -34.9615 | -25.8226
15.9077 | -34.962 | -25.8306
15.9138 | -34.9626 | -25.8385
15.9199 | -34.9631 | -25.8464
15.926 | -34.9636 | -25.8543
15.9321 | -34.9642 | -25.862
15.9382 | -34.9647 | -25.87
15.9443 | -34.9653 | -25.8779
15.9504 | -34.9658 | -25.8859
15.9565 | -34.9664 | -25.8938
15.9626 | -34.9669 | -25.9019
15.9687 | -34.9675 | -25.9096
15.9749 | -34.968 | -25.9173
15.981 | -34.9685 | -25.9252
15.9871 | -34.9691 | -25.9332
15.9932 | -34.9696 | -25.941
15.9993 | -34.9702 | -25.9491
16.0054 | -34.9708 | -25.9566
16.0115 | -34.9713 | -25.9646
16.0176 | -34.9719 | -25.9726
16.0237 | -34.9724 | -25.9806
16.0298 | -34.973 | -25.9884
16.0359 | -34.9735 | -25.9964
16.042 | -34.9741 | -26.0042
16.0481 | -34.9746 | -26.0123
16.0542 | -34.9752 | -26.0204
16.0603 | -34.9757 | -26.0282
16.0664 | -34.9763 | -26.0363
16.0725 | -34.9768 | -26.0442
16.0786 | -34.9774 | -26.052
16.0847 | -34.978 | -26.0601
16.0908 | -34.9785 | -26.068
16.0969 | -34.9791 | -26.0761
16.103 | -34.9796 | -26.0841
16.1091 | -34.9802 | -26.0918
16.1152 | -34.9807 | -26.0997
16.1214 | -34.9813 | -26.1077
16.1275 | -34.9819 | -26.1157
16.1336 | -34.9824 | -26.1235
16.1397 | -34.983 | -26.1315
16.1458 | -34.9835 | -26.1391
16.1519 | -34.9841 | -26.1472
16.158 | -34.9846 | -26.155
16.1641 | -34.9852 | -26.1629
16.1702 | -34.9858 | -26.1708
16.1763 | -34.9863 | -26.1787
16.1824 | -34.9869 | -26.1864
16.1885 | -34.9874 | -26.1943
16.1946 | -34.988 | -26.202
16.2007 | -34.9885 | -26.2099
16.2068 | -34.9891 | -26.2177
16.2129 | -34.9896 | -26.2257
16.219 | -34.9902 | -26.2333
16.2251 | -34.9907 | -26.2413
16.2312 | -34.9913 | -26.2492
16.2373 | -34.9918 | -26.2568
16.2434 | -34.9924 | -26.2646
16.2495 | -34.993 | -26.2723
16.2556 | -34.9935 | -26.2804
16.2618 | -34.9941 | -26.2881
16.2679 | -34.9946 | -26.2961
16.274 | -34.9952 | -26.3039
16.2801 | -34.9957 | -26.3118
16.2862 | -34.9963 | -26.3195
16.2923 | -34.9969 | -26.3275
16.2984 | -34.9974 | -26.3353
16.3045 | -34.9979 | -26.3432
16.3106 | -34.9985 | -26.3511
16.3167 | -34.999 | -26.3589
16.3228 | -34.9996 | -26.3666
16.3289 | -35.0001 | -26.3744
16.335 | -35.0007 | -26.3822
16.3411 | -35.0012 | -26.3899
16.3472 | -35.0018 | -26.3979
16.3533 | -35.0023 | -26.4058
16.3594 | -35.0029 | -26.4132
16.3655 | -35.0034 | -26.4212
16.3716 | -35.004 | -26.4288
16.3777 | -35.0046 | -26.4366
16.3838 | -35.0051 | -26.4443
16.3899 | -35.0057 | -26.4519
16.396 | -35.0062 | -26.4598
16.4021 | -35.0068 | -26.4676
16.4083 | -35.0073 | -26.4754
16.4144 | -35.0079 | -26.4834
16.4205 | -35.0084 | -26.4912
16.4266 | -35.009 | -26.4987
16.4327 | -35.0095 | -26.5066
16.4388 | -35.0101 | -26.5144
16.4449 | -35.0106 | -26.5221
16.451 | -35.0112 | -26.5301
16.4571 | -35.0117 | -26.5379
16.4632 | -35.0123 | -26.5453
16.4693 | -35.0128 | -26.5531
16.4754 | -35.0134 | -26.561
16.4815 | -35.0139 | -26.5687
16.4876 | -35.0145 | -26.5763
16.4937 | -35.015 | -26.5839
16.4998 | -35.0156 | -26.5913
16.5059 | -35.0161 | -26.5991
16.512 | -35.0167 | -26.607
16.5181 | -35.0173 | -26.615
16.5242 | -35.0178 | -26.6228
16.5303 | -35.0184 | -26.6303
16.5364 | -35.0189 | -26.6379
16.5425 | -35.0195 | -26.6457
16.5487 | -35.02 | -26.6533
16.5548 | -35.0206 | -26.661
16.5609 | -35.0211 | -26.6688
16.567 | -35.0217 | -26.6763
16.5731 | -35.0223 | -26.6841
16.5792 | -35.0228 | -26.6919
16.5853 | -35.0233 | -26.6999
16.5914 | -35.0239 | -26.7077
16.5975 | -35.0244 | -26.7155
16.6036 | -35.025 | -26.7231
16.6097 | -35.0255 | -26.731
16.6158 | -35.0261 | -26.7388
16.6219 | -35.0266 | -26.7465
16.628 | -35.0272 | -26.7543
16.6341 | -35.0277 | -26.762
16.6402 | -35.0283 | -26.7695
16.6463 | -35.0288 | -26.7773
16.6524 | -35.0293 | -26.785
16.6585 | -35.0299 | -26.7928
16.6646 | -35.0305 | -26.8004
16.6707 | -35.031 | -26.808
16.6768 | -35.0316 | -26.8157
16.6829 | -35.0321 | -26.8232
16.689 | -35.0327 | -26.831
16.6952 | -35.0332 | -26.8385
16.7013 | -35.0338 | -26.8461
16.7074 | -35.0344 | -26.8534
16.7135 | -35.0349 | -26.8613
16.7196 | -35.0355 | -26.869
16.7257 | -35.036 | -26.8766
16.7318 | -35.0366 | -26.8842
16.7379 | -35.0371 | -26.8917
16.744 | -35.0377 | -26.8991
16.7501 | -35.0382 | -26.9068
16.7562 | -35.0388 | -26.9144
16.7623 | -35.0393 | -26.922
16.7684 | -35.0399 | -26.9299
16.7745 | -35.0404 | -26.9377
16.7806 | -35.041 | -26.9451
16.7867 | -35.0415 | -26.9528
16.7928 | -35.0421 | -26.9605
16.7989 | -35.0426 | -26.9681
16.805 | -35.0432 | -26.9758
16.8111 | -35.0437 | -26.9833
16.8172 | -35.0443 | -26.9912
16.8233 | -35.0449 | -26.999
16.8294 | -35.0454 | -27.0067
16.8356 | -35.046 | -27.0144
16.8417 | -35.0465 | -27.0219
16.8478 | -35.0471 | -27.0296
16.8539 | -35.0476 | -27.0371
16.86 | -35.0481 | -27.0449
16.8661 | -35.0487 | -27.0525
16.8722 | -35.0492 | -27.0603
16.8783 | -35.0498 | -27.068
16.8844 | -35.0504 | -27.0756
16.8905 | -35.0509 | -27.0833
16.8966 | -35.0514 | -27.091
16.9027 | -35.052 | -27.0986
16.9088 | -35.0525 | -27.1064
16.9149 | -35.0531 | -27.1141
16.921 | -35.0537 | -27.1216
16.9271 | -35.0542 | -27.1295
16.9332 | -35.0548 | -27.1374
16.9393 | -35.0553 | -27.1451
16.9454 | -35.0558 | -27.1527
16.9515 | -35.0563 | -27.1603
16.9576 | -35.0569 | -27.1676
16.9637 | -35.0575 | -27.1751
16.9698 | -35.058 | -27.1827
16.9759 | -35.0586 | -27.1903
16.9821 | -35.0591 | -27.1982
16.9882 | -35.0597 | -27.2056
16.9943 | -35.0602 | -27.2132
17.0004 | -35.0608 | -27.2208
17.0065 | -35.0613 | -27.2285
17.0126 | -35.0618 | -27.236
17.0187 | -35.0624 | -27.2438
17.0248 | -35.0629 | -27.2514
17.0309 | -35.0635 | -27.2589
17.037 | -35.064 | -27.2665
17.0431 | -35.0645 | -27.2741
17.0492 | -35.0651 | -27.2817
17.0553 | -35.0656 | -27.2893
17.0614 | -35.0661 | -27.2969
17.0675 | -35.0667 | -27.3045
17.0736 | -35.0672 | -27.3121
17.0797 | -35.0678 | -27.3196
17.0858 | -35.0683 | -27.3274
17.0919 | -35.0689 | -27.335
17.098 | -35.0694 | -27.3423
17.1041 | -35.07 | -27.3499
17.1102 | -35.0706 | -27.3576
17.1163 | -35.0711 | -27.3653
17.1225 | -35.0716 | -27.3727
17.1286 | -35.0722 | -27.3801
17.1347 | -35.0728 | -27.3877
17.1408 | -35.0733 | -27.3953
17.1469 | -35.0739 | -27.4028
17.153 | -35.0744 | -27.4103
17.1591 | -35.075 | -27.4179
17.1652 | -35.0755 | -27.4252
17.1713 | -35.0761 | -27.4326
17.1774 | -35.0766 | -27.4403
17.1835 | -35.0771 | -27.4478
17.1896 | -35.0776 | -27.4554
17.1957 | -35.0782 | -27.463
17.2018 | -35.0788 | -27.4703
17.2079 | -35.0793 | -27.4778
17.214 | -35.0799 | -27.4855
17.2201 | -35.0804 | -27.4931
17.2262 | -35.0809 | -27.5007
17.2323 | -35.0815 | -27.5082
17.2384 | -35.0821 | -27.5156
17.2445 | -35.0826 | -27.5231
17.2506 | -35.0832 | -27.5307
17.2567 | -35.0837 | -27.5386
17.2628 | -35.0843 | -27.5461
17.269 | -35.0848 | -27.5535
17.2751 | -35.0854 | -27.5609
17.2812 | -35.0859 | -27.5684
17.2873 | -35.0865 | -27.5761
17.2934 | -35.087 | -27.5838
17.2995 | -35.0876 | -27.5914
17.3056 | -35.0881 | -27.5987
17.3117 | -35.0887 | -27.6063
17.3178 | -35.0892 | -27.6136
17.3239 | -35.0898 | -27.621
17.33 | -35.0903 | -27.6286
17.3361 | -35.0908 | -27.636
17.3422 | -35.0913 | -27.6432
17.3483 | -35.0919 | -27.6508
17.3544 | -35.0924 | -27.6583
17.3605 | -35.093 | -27.6658
17.3666 | -35.0935 | -27.6734
17.3727 | -35.0941 | -27.6809
17.3788 | -35.0946 | -27.688
17.3849 | -35.0952 | -27.6955
17.391 | -35.0957 | -27.7031
17.3971 | -35.0963 | -27.7106
17.4032 | -35.0968 | -27.7181
17.4094 | -35.0974 | -27.7255
17.4155 | -35.098 | -27.7328
17.4216 | -35.0985 | -27.7403
17.4277 | -35.0991 | -27.7479
17.4338 | -35.0996 | -27.7554
17.4399 | -35.1002 | -27.763
17.446 | -35.1007 | -27.7702
17.4521 | -35.1012 | -27.7776
17.4582 | -35.1018 | -27.7849
17.4643 | -35.1023 | -27.7924
17.4704 | -35.1028 | -27.7999
17.4765 | -35.1034 | -27.8074
17.4826 | -35.1039 | -27.8147
17.4887 | -35.1045 | -27.8224
17.4948 | -35.105 | -27.8298
17.5009 | -35.1055 | -27.8374
17.507 | -35.1061 | -27.8449
17.5131 | -35.1066 | -27.8524
17.5192 | -35.1072 | -27.8597
17.5253 | -35.1077 | -27.8671
17.5314 | -35.1082 | -27.8747
17.5375 | -35.1088 | -27.8822
17.5436 | -35.1093 | -27.8898
17.5497 | -35.1099 | -27.8972
17.5559 | -35.1104 | -27.9046
17.562 | -35.111 | -27.9119
17.5681 | -35.1115 | -27.9193
17.5742 | -35.1121 | -27.9267
17.5803 | -35.1126 | -27.9342
17.5864 | -35.1132 | -27.9416
17.5925 | -35.1138 | -27.9491
17.5986 | -35.1144 | -27.9566
17.6047 | -35.1149 | -27.964
17.6108 | -35.1154 | -27.9714
17.6169 | -35.116 | -27.9788
17.623 | -35.1166 | -27.986
17.6291 | -35.1171 | -27.9934
17.6352 | -35.1176 | -28.0009
17.6413 | -35.1182 | -28.0084
17.6474 | -35.1187 | -28.016
17.6535 | -35.1193 | -28.0234
17.6596 | -35.1198 | -28.0308
17.6657 | -35.1204 | -28.0382
17.6718 | -35.1209 | -28.0457
17.6779 | -35.1215 | -28.0531
17.684 | -35.122 | -28.0606
17.6901 | -35.1225 | -28.0682
17.6963 | -35.1231 | -28.0754
17.7024 | -35.1237 | -28.0829
17.7085 | -35.1242 | -28.0903
17.7146 | -35.1248 | -28.0978
17.7207 | -35.1253 | -28.1051
17.7268 | -35.1259 | -28.1124
17.7329 | -35.1264 | -28.1199
17.739 | -35.127 | -28.1274
17.7451 | -35.1275 | -28.1349
17.7512 | -35.1281 | -28.1424
17.7573 | -35.1287 | -28.1498
17.7634 | -35.1292 | -28.157
17.7695 | -35.1298 | -28.1645
17.7756 | -35.1303 | -28.172
17.7817 | -35.1309 | -28.1794
17.7878 | -35.1314 | -28.1869
17.7939 | -35.132 | -28.1944
17.8 | -35.1325 | -28.2018
17.8061 | -35.1331 | -28.2093
17.8122 | -35.1336 | -28.2167
17.8183 | -35.1342 | -28.2241
17.8244 | -35.1347 | -28.2315
17.8305 | -35.1353 | -28.2389
17.8366 | -35.1358 | -28.2461
17.8428 | -35.1364 | -28.2535
17.8489 | -35.137 | -28.2607
17.855 | -35.1375 | -28.2681
17.8611 | -35.1381 | -28.2753
17.8672 | -35.1386 | -28.2828
17.8733 | -35.1392 | -28.2903
17.8794 | -35.1398 | -28.2977
17.8855 | -35.1403 | -28.305
17.8916 | -35.1408 | -28.3126
17.8977 | -35.1414 | -28.3201
17.9038 | -35.1419 | -28.3273
17.9099 | -35.1425 | -28.3346
17.916 | -35.143 | -28.342
17.9221 | -35.1436 | -28.3496
17.9282 | -35.1441 | -28.3571
17.9343 | -35.1447 | -28.3645
17.9404 | -35.1452 | -28.3717
17.9465 | -35.1458 | -28.3791
17.9526 | -35.1463 | -28.3867
17.9587 | -35.1469 | -28.394
17.9648 | -35.1474 | -28.4014
17.9709 | -35.148 | -28.4088
17.977 | -35.1485 | -28.4162
17.9832 | -35.149 | -28.4236
17.9893 | -35.1496 | -28.431
17.9954 | -35.1501 | -28.4385
18.0015 | -35.1506 | -28.4459
18.0076 | -35.1512 | -28.4532
18.0137 | -35.1517 | -28.4606
18.0198 | -35.1523 | -28.4678
18.0259 | -35.1528 | -28.4753
18.032 | -35.1533 | -28.4827
18.0381 | -35.1539 | -28.4902
18.0442 | -35.1544 | -28.4975
18.0503 | -35.155 | -28.5049
18.0564 | -35.1555 | -28.5124
18.0625 | -35.1561 | -28.5199
18.0686 | -35.1566 | -28.5273
18.0747 | -35.1572 | -28.5347
18.0808 | -35.1577 | -28.5419
18.0869 | -35.1583 | -28.5493
18.093 | -35.1588 | -28.5567
18.0991 | -35.1593 | -28.5641
18.1052 | -35.1599 | -28.5716
18.1113 | -35.1605 | -28.579
18.1174 | -35.161 | -28.5864
18.1236 | -35.1616 | -28.5939
18.1297 | -35.1621 | -28.6014
18.1358 | -35.1627 | -28.6087
18.1419 | -35.1632 | -28.616
18.148 | -35.1638 | -28.6233
18.1541 | -35.1643 | -28.6304
18.1602 | -35.1649 | -28.6378
18.1663 | -35.1654 | -28.645
18.1724 | -35.166 | -28.6524
18.1785 | -35.1665 | -28.6595
18.1846 | -35.167 | -28.6668
18.1907 | -35.1676 | -28.6742
18.1968 | -35.1681 | -28.6816
18.2029 | -35.1687 | -28.6891
18.209 | -35.1692 | -28.6965
18.2151 | -35.1697 | -28.704
18.2212 | -35.1703 | -28.7113
18.2273 | -35.1709 | -28.7185
18.2334 | -35.1714 | -28.7259
18.2395 | -35.1719 | -28.7333
18.2456 | -35.1725 | -28.7405
18.2517 | -35.1731 | -28.748
18.2578 | -35.1736 | -28.7551
18.2639 | -35.1742 | -28.7625
18.2701 | -35.1747 | -28.7699
18.2762 | -35.1752 | -28.7772
18.2823 | -35.1758 | -28.7845
18.2884 | -35.1763 | -28.7918
18.2945 | -35.1769 | -28.7988
18.3006 | -35.1774 | -28.8064
18.3067 | -35.178 | -28.8137
18.3128 | -35.1785 | -28.821
18.3189 | -35.1791 | -28.8284
18.325 | -35.1796 | -28.8357
18.3311 | -35.1801 | -28.8431
18.3372 | -35.1807 | -28.8504
18.3433 | -35.1813 | -28.8577
18.3494 | -35.1818 | -28.865
18.3555 | -35.1823 | -28.8724
18.3616 | -35.1829 | -28.8795
18.3677 | -35.1834 | -28.8868
18.3738 | -35.184 | -28.8941
18.3799 | -35.1845 | -28.9015
18.386 | -35.1851 | -28.9089
18.3921 | -35.1856 | -28.9162
18.3982 | -35.1862 | -28.9234
18.4043 | -35.1867 | -28.9308
18.4105 | -35.1873 | -28.9383
18.4166 | -35.1878 | -28.9454
18.4227 | -35.1884 | -28.9527
18.4288 | -35.1889 | -28.9601
18.4349 | -35.1894 | -28.9674
18.441 | -35.19 | -28.9745
18.4471 | -35.1905 | -28.9817
18.4532 | -35.1911 | -28.9891
18.4593 | -35.1916 | -28.9964
18.4654 | -35.1921 | -29.0036
18.4715 | -35.1927 | -29.0111
18.4776 | -35.1932 | -29.0187
18.4837 | -35.1937 | -29.0261
18.4898 | -35.1943 | -29.0335
18.4959 | -35.1948 | -29.041
18.502 | -35.1953 | -29.0482
18.5081 | -35.1959 | -29.0556
18.5142 | -35.1964 | -29.063
18.5203 | -35.1969 | -29.0703
18.5264 | -35.1975 | -29.0778
18.5325 | -35.198 | -29.0849
18.5386 | -35.1986 | -29.0924
18.5447 | -35.1991 | -29.0996
18.5508 | -35.1997 | -29.1068
18.557 | -35.2002 | -29.1143
18.5631 | -35.2007 | -29.1217
18.5692 | -35.2013 | -29.1289
18.5753 | -35.2018 | -29.1364
18.5814 | -35.2024 | -29.1435
18.5875 | -35.2029 | -29.1509
18.5936 | -35.2035 | -29.1582
18.5997 | -35.204 | -29.1656
18.6058 | -35.2046 | -29.1726
18.6119 | -35.2051 | -29.1799
18.618 | -35.2056 | -29.1874
18.6241 | -35.2062 | -29.1951
18.6302 | -35.2067 | -29.2026
18.6363 | -35.2073 | -29.2097
18.6424 | -35.2078 | -29.217
18.6485 | -35.2083 | -29.2244
18.6546 | -35.2089 | -29.2317
18.6607 | -35.2094 | -29.239
18.6668 | -35.21 | -29.2463
18.6729 | -35.2105 | -29.2537
18.679 | -35.211 | -29.2607
18.6851 | -35.2116 | -29.268
18.6912 | -35.2121 | -29.2753
18.6974 | -35.2126 | -29.2827
18.7035 | -35.2132 | -29.2901
18.7096 | -35.2137 | -29.2975
18.7157 | -35.2143 | -29.3047
18.7218 | -35.2148 | -29.3119
18.7279 | -35.2154 | -29.3193
18.734 | -35.2159 | -29.3268
18.7401 | -35.2165 | -29.3341
18.7462 | -35.217 | -29.3413
18.7523 | -35.2176 | -29.3485
18.7584 | -35.2181 | -29.3561
18.7645 | -35.2187 | -29.3633
18.7706 | -35.2192 | -29.3706
18.7767 | -35.2198 | -29.3779
18.7828 | -35.2203 | -29.3851
18.7889 | -35.2208 | -29.3926
18.795 | -35.2214 | -29.4
18.8011 | -35.2219 | -29.4074
18.8072 | -35.2225 | -29.4148
18.8133 | -35.223 | -29.4222
18.8194 | -35.2235 | -29.4295
18.8255 | -35.2241 | -29.4367
18.8316 | -35.2246 | -29.4441
18.8377 | -35.2252 | -29.4513
18.8439 | -35.2257 | -29.4584
18.85 | -35.2263 | -29.4658
18.8561 | -35.2268 | -29.473
18.8622 | -35.2273 | -29.4804
18.8683 | -35.2279 | -29.4874
18.8744 | -35.2284 | -29.4948
18.8805 | -35.229 | -29.5023
18.8866 | -35.2295 | -29.5095
18.8927 | -35.2301 | -29.5167
18.8988 | -35.2306 | -29.5239
18.9049 | -35.2312 | -29.5311
18.911 | -35.2317 | -29.5383
18.9171 | -35.2323 | -29.5457
18.9232 | -35.2328 | -29.5527
18.9293 | -35.2334 | -29.5599
18.9354 | -35.2339 | -29.5673
18.9415 | -35.2345 | -29.5746
18.9476 | -35.235 | -29.5817
18.9537 | -35.2355 | -29.5891
18.9598 | -35.236 | -29.5963
18.9659 | -35.2366 | -29.6035
18.972 | -35.2371 | -29.6107
18.9781 | -35.2376 | -29.618
18.9843 | -35.2381 | -29.6253
18.9904 | -35.2387 | -29.6326
18.9965 | -35.2393 | -29.6397
19.0026 | -35.2398 | -29.6471
19.0087 | -35.2404 | -29.6545
19.0148 | -35.2409 | -29.6617
19.0209 | -35.2415 | -29.6691
19.027 | -35.242 | -29.6766
19.0331 | -35.2426 | -29.6837
19.0392 | -35.2432 | -29.6911
19.0453 | -35.2437 | -29.6983
19.0514 | -35.2443 | -29.7055
19.0575 | -35.2448 | -29.7129
19.0636 | -35.2454 | -29.7199
19.0697 | -35.2459 | -29.7271
19.0758 | -35.2464 | -29.7344
19.0819 | -35.247 | -29.7416
19.088 | -35.2475 | -29.7489
19.0941 | -35.248 | -29.7563
19.1002 | -35.2486 | -29.7632
19.1063 | -35.2491 | -29.7704
19.1124 | -35.2496 | -29.7778
19.1185 | -35.2501 | -29.7851
19.1246 | -35.2507 | -29.7924
19.1308 | -35.2512 | -29.7999
19.1369 | -35.2518 | -29.8071
19.143 | -35.2523 | -29.8143
19.1491 | -35.2529 | -29.8218
19.1552 | -35.2534 | -29.829
19.1613 | -35.2539 | -29.8363
19.1674 | -35.2545 | -29.8437
19.1735 | -35.255 | -29.8509
19.1796 | -35.2556 | -29.8582
19.1857 | -35.2561 | -29.8656
19.1918 | -35.2567 | -29.873
19.1979 | -35.2573 | -29.8804
19.204 | -35.2579 | -29.8877
19.2101 | -35.2584 | -29.8949
19.2162 | -35.2589 | -29.902
19.2223 | -35.2595 | -29.9092
19.2284 | -35.26 | -29.9164
19.2345 | -35.2605 | -29.9236
19.2406 | -35.2611 | -29.9305
19.2467 | -35.2616 | -29.9379
19.2528 | -35.2622 | -29.9452
19.2589 | -35.2627 | -29.9524
19.265 | -35.2633 | -29.9597
19.2712 | -35.2638 | -29.967
19.2773 | -35.2643 | -29.9739
19.2834 | -35.2649 | -29.981
19.2895 | -35.2654 | -29.9883
19.2956 | -35.2659 | -29.9957
19.3017 | -35.2665 | -30.0029
19.3078 | -35.267 | -30.0102
19.3139 | -35.2676 | -30.0169
19.32 | -35.2681 | -30.024
19.3261 | -35.2686 | -30.0312
19.3322 | -35.2692 | -30.0383
19.3383 | -35.2698 | -30.0452
19.3444 | -35.2703 | -30.0523
19.3505 | -35.2708 | -30.0594
19.3566 | -35.2714 | -30.0666
19.3627 | -35.2719 | -30.0737
19.3688 | -35.2725 | -30.081
19.3749 | -35.2731 | -30.0882
19.381 | -35.2736 | -30.0953
19.3871 | -35.2741 | -30.1024
19.3932 | -35.2747 | -30.1096
19.3993 | -35.2752 | -30.1167
19.4054 | -35.2757 | -30.1243
19.4115 | -35.2762 | -30.1316
19.4177 | -35.2767 | -30.139
19.4238 | -35.2773 | -30.1463
19.4299 | -35.2778 | -30.1537
19.436 | -35.2783 | -30.161
19.4421 | -35.2788 | -30.168
19.4482 | -35.2793 | -30.1752
19.4543 | -35.2799 | -30.1819
19.4604 | -35.2805 | -30.1892
19.4665 | -35.281 | -30.1962
19.4726 | -35.2815 | -30.2035
19.4787 | -35.2821 | -30.2107
19.4848 | -35.2826 | -30.2178
19.4909 | -35.2831 | -30.2246
19.497 | -35.2837 | -30.2321
19.5031 | -35.2842 | -30.2392
19.5092 | -35.2847 | -30.2464
19.5153 | -35.2852 | -30.2537
19.5214 | -35.2857 | -30.2608
19.5275 | -35.2862 | -30.2681
19.5336 | -35.2867 | -30.2756
19.5397 | -35.2873 | -30.2829
19.5458 | -35.2878 | -30.2902
19.5519 | -35.2883 | -30.2975
19.5581 | -35.2889 | -30.3044
19.5642 | -35.2894 | -30.3118
19.5703 | -35.2899 | -30.3186
19.5764 | -35.2904 | -30.3256
19.5825 | -35.291 | -30.3332
19.5886 | -35.2915 | -30.3404
19.5947 | -35.292 | -30.3468
19.6008 | -35.2926 | -30.3539
19.6069 | -35.2931 | -30.3611
19.613 | -35.2937 | -30.3684
19.6191 | -35.2943 | -30.3757
19.6252 | -35.2949 | -30.3827
19.6313 | -35.2954 | -30.3899
19.6374 | -35.2959 | -30.3974
19.6435 | -35.2965 | -30.4046
19.6496 | -35.297 | -30.4117
19.6557 | -35.2975 | -30.4191
19.6618 | -35.298 | -30.4257
19.6679 | -35.2985 | -30.4331
19.674 | -35.2991 | -30.44
19.6801 | -35.2996 | -30.4472
19.6862 | -35.3002 | -30.4543
19.6923 | -35.3008 | -30.4611
19.6984 | -35.3013 | -30.4681
19.7046 | -35.3018 | -30.4752
19.7107 | -35.3024 | -30.4821
19.7168 | -35.3029 | -30.4893
19.7229 | -35.3034 | -30.4964
19.729 | -35.3039 | -30.5034
19.7351 | -35.3044 | -30.5106
19.7412 | -35.305 | -30.5179
19.7473 | -35.3055 | -30.525
19.7534 | -35.3061 | -30.5323
19.7595 | -35.3066 | -30.5393
19.7656 | -35.3072 | -30.5468
19.7717 | -35.3077 | -30.5534
19.7778 | -35.3083 | -30.5602
19.7839 | -35.3088 | -30.5672
19.79 | -35.3093 | -30.5743
19.7961 | -35.3099 | -30.5815
19.8022 | -35.3104 | -30.5882
19.8083 | -35.3109 | -30.5951
19.8144 | -35.3115 | -30.6023
19.8205 | -35.312 | -30.6093
19.8266 | -35.3126 | -30.6162
19.8327 | -35.3131 | -30.6232
19.8388 | -35.3137 | -30.6307
19.845 | -35.3142 | -30.6379
19.8511 | -35.3148 | -30.645
19.8572 | -35.3153 | -30.652
19.8633 | -35.3159 | -30.659
19.8694 | -35.3164 | -30.6666
19.8755 | -35.3169 | -30.6734
19.8816 | -35.3174 | -30.6805
19.8877 | -35.318 | -30.6879
19.8938 | -35.3186 | -30.6951
19.8999 | -35.3191 | -30.7019
19.906 | -35.3196 | -30.7089
19.9121 | -35.3201 | -30.7157
19.9182 | -35.3207 | -30.7229
19.9243 | -35.3212 | -30.7299
19.9304 | -35.3217 | -30.7368
19.9365 | -35.3223 | -30.7438
19.9426 | -35.3228 | -30.7507
19.9487 | -35.3234 | -30.7579
19.9548 | -35.3239 | -30.7651
19.9609 | -35.3245 | -30.7721
19.967 | -35.325 | -30.7794
19.9731 | -35.3255 | -30.7864
19.9792 | -35.3261 | -30.793
19.9853 | -35.3266 | -30.8
19.9915 | -35.3271 | -30.8067
19.9976 | -35.3276 | -30.8135
20.0037 | -35.3282 | -30.8205
20.0098 | -35.3287 | -30.8273
20.0159 | -35.3293 | -30.834
20.022 | -35.3298 | -30.841
20.0281 | -35.3304 | -30.8482
20.0342 | -35.3309 | -30.8552
20.0403 | -35.3314 | -30.8623
20.0464 | -35.3319 | -30.8693
20.0525 | -35.3324 | -30.8763
20.0586 | -35.333 | -30.8836
20.0647 | -35.3335 | -30.8906
20.0708 | -35.3341 | -30.8979
20.0769 | -35.3346 | -30.9048
20.083 | -35.3351 | -30.9114
20.0891 | -35.3356 | -30.9186
20.0952 | -35.3361 | -30.9255
20.1013 | -35.3367 | -30.9328
20.1074 | -35.3372 | -30.9397
20.1135 | -35.3378 | -30.9466
20.1196 | -35.3383 | -30.9537
20.1257 | -35.3388 | -30.9607
20.1319 | -35.3393 | -30.9679
20.138 | -35.3398 | -30.9751
20.1441 | -35.3404 | -30.9818
20.1502 | -35.3409 | -30.9895
20.1563 | -35.3414 | -30.9967
20.1624 | -35.3419 | -31.0036
20.1685 | -35.3424 | -31.0107
20.1746 | -35.343 | -31.0177
20.1807 | -35.3435 | -31.0242
20.1868 | -35.3441 | -31.0313
20.1929 | -35.3446 | -31.038
20.199 | -35.3451 | -31.0452
20.2051 | -35.3456 | -31.052
20.2112 | -35.3462 | -31.0592
20.2173 | -35.3467 | -31.0663
20.2234 | -35.3472 | -31.0733
20.2295 | -35.3478 | -31.0804
20.2356 | -35.3483 | -31.087
20.2417 | -35.3488 | -31.0939
20.2478 | -35.3493 | -31.1006
20.2539 | -35.3499 | -31.108
20.26 | -35.3505 | -31.1149
20.2661 | -35.351 | -31.1222
20.2723 | -35.3516 | -31.1289
20.2784 | -35.3521 | -31.1356
20.2845 | -35.3527 | -31.1423
20.2906 | -35.3533 | -31.1489
20.2967 | -35.3537 | -31.1558
20.3028 | -35.3542 | -31.1626
20.3089 | -35.3548 | -31.1697
20.315 | -35.3553 | -31.1765
20.3211 | -35.3559 | -31.1837
20.3272 | -35.3564 | -31.1906
20.3333 | -35.3569 | -31.1978
20.3394 | -35.3574 | -31.2048
20.3455 | -35.358 | -31.212
20.3516 | -35.3585 | -31.2188
20.3577 | -35.359 | -31.2255
20.3638 | -35.3595 | -31.2325
20.3699 | -35.36 | -31.2394
20.376 | -35.3605 | -31.2466
20.3821 | -35.3611 | -31.2535
20.3882 | -35.3616 | -31.2607
20.3943 | -35.3621 | -31.2678
20.4004 | -35.3626 | -31.2743
20.4065 | -35.3631 | -31.2813
20.4126 | -35.3637 | -31.288
20.4188 | -35.3642 | -31.2953
20.4249 | -35.3647 | -31.3022
20.431 | -35.3652 | -31.3093
20.4371 | -35.3656 | -31.3161
20.4432 | -35.3662 | -31.3227
20.4493 | -35.3667 | -31.3294
20.4554 | -35.3672 | -31.3366
20.4615 | -35.3677 | -31.3435
20.4676 | -35.3682 | -31.3511
20.4737 | -35.3687 | -31.3579
20.4798 | -35.3692 | -31.3647
20.4859 | -35.3697 | -31.3719
20.492 | -35.3703 | -31.3787
20.4981 | -35.3708 | -31.3859
20.5042 | -35.3713 | -31.3926
20.5103 | -35.3718 | -31.3992
20.5164 | -35.3724 | -31.4058
20.5225 | -35.3728 | -31.4129
20.5286 | -35.3733 | -31.42
20.5347 | -35.374 | -31.427
20.5408 | -35.3745 | -31.4336
20.5469 | -35.375 | -31.4403
20.553 | -35.3756 | -31.4474
20.5592 | -35.3761 | -31.4541
20.5653 | -35.3767 | -31.4607
20.5714 | -35.3772 | -31.4677
20.5775 | -35.3777 | -31.4742
20.5836 | -35.3782 | -31.4814
20.5897 | -35.3788 | -31.4884
20.5958 | -35.3793 | -31.4957
20.6019 | -35.3798 | -31.5027
20.608 | -35.3803 | -31.51
20.6141 | -35.3809 | -31.5162
20.6202 | -35.3814 | -31.5234
20.6263 | -35.382 | -31.5302
20.6324 | -35.3825 | -31.5369
20.6385 | -35.383 | -31.5438
20.6446 | -35.3836 | -31.5508
20.6507 | -35.3842 | -31.5576
20.6568 | -35.3847 | -31.5646
20.6629 | -35.3853 | -31.5716
20.669 | -35.3859 | -31.5783
20.6751 | -35.3864 | -31.5853
20.6812 | -35.3869 | -31.5916
20.6873 | -35.3874 | -31.5987
20.6934 | -35.3879 | -31.605
20.6995 | -35.3885 | -31.6117
20.7057 | -35.3891 | -31.6186
20.7118 | -35.3896 | -31.6255
20.7179 | -35.3902 | -31.6322
20.724 | -35.3907 | -31.6391
20.7301 | -35.3912 | -31.6457
20.7362 | -35.3918 | -31.6525
20.7423 | -35.3923 | -31.6594
20.7484 | -35.3929 | -31.6663
20.7545 | -35.3934 | -31.6726
20.7606 | -35.394 | -31.6795
20.7667 | -35.3945 | -31.6863
20.7728 | -35.3951 | -31.6927
20.7789 | -35.3956 | -31.6998
20.785 | -35.3961 | -31.7073
20.7911 | -35.3966 | -31.7138
20.7972 | -35.3972 | -31.7212
20.8033 | -35.3977 | -31.7281
20.8094 | -35.3982 | -31.7353
20.8155 | -35.3988 | -31.7421
20.8216 | -35.3993 | -31.7489
20.8277 | -35.3999 | -31.7557
20.8338 | -35.4004 | -31.7625
20.8399 | -35.4009 | -31.7691
20.8461 | -35.4014 | -31.776
20.8522 | -35.4019 | -31.783
20.8583 | -35.4025 | -31.7896
20.8644 | -35.403 | -31.7961
20.8705 | -35.4035 | -31.8029
20.8766 | -35.4041 | -31.8096
20.8827 | -35.4046 | -31.8162
20.8888 | -35.4051 | -31.823
20.8949 | -35.4056 | -31.8297
20.901 | -35.4062 | -31.8367
20.9071 | -35.4067 | -31.8436
20.9132 | -35.4073 | -31.8503
20.9193 | -35.4078 | -31.8572
20.9254 | -35.4083 | -31.8639
20.9315 | -35.4089 | -31.8704
20.9376 | -35.4094 | -31.8771
20.9437 | -35.41 | -31.8837
20.9498 | -35.4105 | -31.8901
20.9559 | -35.411 | -31.8968
20.962 | -35.4115 | -31.9036
20.9681 | -35.4121 | -31.9099
20.9742 | -35.4126 | -31.9168
20.9803 | -35.4132 | -31.9236
20.9864 | -35.4137 | -31.9305
20.9926 | -35.4143 | -31.9372
20.9987 | -35.4148 | -31.9436
21.0048 | -35.4153 | -31.9502
21.0109 | -35.4158 | -31.9569
21.017 | -35.4164 | -31.9635
21.0231 | -35.4169 | -31.9702
21.0292 | -35.4174 | -31.9768
21.0353 | -35.418 | -31.983
21.0414 | -35.4185 | -31.9896
21.0475 | -35.419 | -31.9966
21.0536 | -35.4195 | -32.0029
21.0597 | -35.4201 | -32.0094
21.0658 | -35.4207 | -32.0163
21.0719 | -35.4211 | -32.0226
21.078 | -35.4217 | -32.0296
21.0841 | -35.4222 | -32.0364
21.0902 | -35.4229 | -32.043
21.0963 | -35.4235 | -32.0506
21.1024 | -35.424 | -32.057
21.1085 | -35.4246 | -32.0637
21.1146 | -35.4251 | -32.0708
21.1207 | -35.4257 | -32.0777
21.1268 | -35.4262 | -32.0846
21.133 | -35.4267 | -32.0912
21.1391 | -35.4273 | -32.0976
21.1452 | -35.4278 | -32.1039
21.1513 | -35.4283 | -32.1106
21.1574 | -35.4288 | -32.1173
21.1635 | -35.4293 | -32.1238
21.1696 | -35.4299 | -32.1305
21.1757 | -35.4305 | -32.1364
21.1818 | -35.4309 | -32.1431
21.1879 | -35.4315 | -32.1501
21.194 | -35.4321 | -32.1568
21.2001 | -35.4326 | -32.1637
21.2062 | -35.4333 | -32.1709
21.2123 | -35.434 | -32.1777
21.2184 | -35.4344 | -32.1847
21.2245 | -35.4347 | -32.189
21.2306 | -35.4353 | -32.1955
21.2367 | -35.4358 | -32.2028
21.2428 | -35.4363 | -32.2096
21.2489 | -35.4368 | -32.216
21.255 | -35.4374 | -32.2226
21.2611 | -35.4379 | -32.2291
21.2672 | -35.4384 | -32.2364
21.2733 | -35.4389 | -32.2429
21.2795 | -35.4395 | -32.2487
21.2856 | -35.44 | -32.255
21.2917 | -35.4406 | -32.262
21.2978 | -35.4411 | -32.2682
21.3039 | -35.4417 | -32.2748
21.31 | -35.4422 | -32.2812
21.3161 | -35.4428 | -32.2869
21.3222 | -35.4433 | -32.2931
21.3283 | -35.4438 | -32.2995
21.3344 | -35.4444 | -32.3059
21.3405 | -35.4449 | -32.3126
21.3466 | -35.4454 | -32.3192
21.3527 | -35.446 | -32.3253
21.3588 | -35.4466 | -32.3317
21.3649 | -35.4471 | -32.3389
21.371 | -35.4476 | -32.3453
21.3771 | -35.4482 | -32.352
21.3832 | -35.4487 | -32.3583
21.3893 | -35.4493 | -32.3647
21.3954 | -35.4498 | -32.3717
21.4015 | -35.4504 | -32.3782
21.4076 | -35.4509 | -32.3848
21.4137 | -35.4515 | -32.3912
21.4199 | -35.452 | -32.3978
21.426 | -35.4525 | -32.4036
21.4321 | -35.453 | -32.4105
21.4382 | -35.4535 | -32.4175
21.4443 | -35.454 | -32.424
21.4504 | -35.4545 | -32.4304
21.4565 | -35.455 | -32.4363
21.4626 | -35.4555 | -32.4433
21.4687 | -35.456 | -32.4501
21.4748 | -35.4565 | -32.4566
21.4809 | -35.457 | -32.4627
21.487 | -35.4575 | -32.4693
21.4931 | -35.4581 | -32.4755
21.4992 | -35.4587 | -32.482
21.5053 | -35.4593 | -32.4883
21.5114 | -35.4598 | -32.4948
21.5175 | -35.4603 | -32.5012
21.5236 | -35.4608 | -32.5078
21.5297 | -35.4614 | -32.5137
21.5358 | -35.462 | -32.5203
21.5419 | -35.4626 | -32.5271
21.548 | -35.4631 | -32.5341
21.5541 | -35.4636 | -32.5402
21.5602 | -35.4641 | -32.5468
21.5664 | -35.4646 | -32.5529
21.5725 | -35.4652 | -32.5597
21.5786 | -35.4657 | -32.5661
21.5847 | -35.4662 | -32.5729
21.5908 | -35.4667 | -32.5794
21.5969 | -35.4673 | -32.5854
21.603 | -35.4678 | -32.5919
21.6091 | -35.4683 | -32.5987
21.6152 | -35.4689 | -32.605
21.6213 | -35.4694 | -32.6116
21.6274 | -35.4698 | -32.6181
21.6335 | -35.4703 | -32.6239
21.6396 | -35.4708 | -32.6306
21.6457 | -35.4713 | -32.6375
21.6518 | -35.4719 | -32.644
21.6579 | -35.4724 | -32.6501
21.664 | -35.4729 | -32.6571
21.6701 | -35.4734 | -32.6635
21.6762 | -35.4739 | -32.6702
21.6823 | -35.4744 | -32.6767
21.6884 | -35.4749 | -32.6833
21.6945 | -35.4755 | -32.6901
21.7006 | -35.476 | -32.6965
21.7068 | -35.4766 | -32.7022
21.7129 | -35.4771 | -32.7084
21.719 | -35.4776 | -32.7147
21.7251 | -35.4781 | -32.7214
21.7312 | -35.4787 | -32.728
21.7373 | -35.4791 | -32.7345
21.7434 | -35.4796 | -32.7414
21.7495 | -35.4801 | -32.7477
21.7556 | -35.4807 | -32.7541
21.7617 | -35.4813 | -32.7609
21.7678 | -35.4819 | -32.7677
21.7739 | -35.4824 | -32.7744
21.78 | -35.483 | -32.781
21.7861 | -35.4835 | -32.7876
21.7922 | -35.484 | -32.7944
21.7983 | -35.4844 | -32.8012
21.8044 | -35.4849 | -32.8074
21.8105 | -35.4854 | -32.8136
21.8166 | -35.4859 | -32.8206
21.8227 | -35.4865 | -32.8273
21.8288 | -35.487 | -32.8338
21.8349 | -35.4875 | -32.8403
21.841 | -35.488 | -32.847
21.8471 | -35.4886 | -32.8534
21.8533 | -35.4891 | -32.8601
21.8594 | -35.4897 | -32.8667
21.8655 | -35.4903 | -32.8732
21.8716 | -35.4907 | -32.8798
21.8777 | -35.4913 | -32.8862
21.8838 | -35.4918 | -32.8926
21.8899 | -35.4923 | -32.8992
21.896 | -35.4928 | -32.9059
21.9021 | -35.4934 | -32.9127
21.9082 | -35.494 | -32.9193
21.9143 | -35.4945 | -32.9246
21.9204 | -35.495 | -32.9307
21.9265 | -35.4956 | -32.9371
21.9326 | -35.4961 | -32.9431
21.9387 | -35.4966 | -32.9493
21.9448 | -35.4972 | -32.9555
21.9509 | -35.4977 | -32.9625
21.957 | -35.4983 | -32.9691
21.9631 | -35.4987 | -32.975
21.9692 | -35.4993 | -32.9818
21.9753 | -35.4999 | -32.9882
21.9814 | -35.5004 | -32.9947
21.9875 | -35.5009 | -33.0013
21.9937 | -35.5015 | -33.0078
21.9998 | -35.5021 | -33.0144
22.0059 | -35.5026 | -33.021
22.012 | -35.5032 | -33.0269
22.0181 | -35.5038 | -33.0334
22.0242 | -35.5044 | -33.0401
22.0303 | -35.5049 | -33.047
22.0364 | -35.5056 | -33.0538
22.0425 | -35.5061 | -33.0603
22.0486 | -35.5067 | -33.067
22.0547 | -35.5072 | -33.0732
22.0608 | -35.5077 | -33.0795
22.0669 | -35.5082 | -33.0866
22.073 | -35.5087 | -33.0926
22.0791 | -35.5093 | -33.0989
22.0852 | -35.5098 | -33.1055
22.0913 | -35.5103 | -33.1121
22.0974 | -35.5108 | -33.1184
22.1035 | -35.5114 | -33.1245
22.1096 | -35.5119 | -33.1309
22.1157 | -35.5124 | -33.1372
22.1218 | -35.5129 | -33.1438
22.1279 | -35.5135 | -33.1498
22.134 | -35.514 | -33.1558
22.1402 | -35.5145 | -33.1617
22.1463 | -35.5151 | -33.168
22.1524 | -35.5156 | -33.1743
22.1585 | -35.5161 | -33.1809
22.1646 | -35.5166 | -33.1876
22.1707 | -35.5171 | -33.1939
22.1768 | -35.5176 | -33.2005
22.1829 | -35.5182 | -33.207
22.189 | -35.5187 | -33.2133
22.1951 | -35.5193 | -33.2197
22.2012 | -35.5198 | -33.226
22.2073 | -35.5204 | -33.2326
22.2134 | -35.5209 | -33.2393
22.2195 | -35.5214 | -33.2458
22.2256 | -35.522 | -33.2527
22.2317 | -35.5225 | -33.2588
22.2378 | -35.523 | -33.265
22.2439 | -35.5235 | -33.2714
22.25 | -35.524 | -33.2777
22.2561 | -35.5245 | -33.2838
22.2622 | -35.5251 | -33.2905
22.2683 | -35.5256 | -33.2964
22.2744 | -35.5261 | -33.3028
22.2806 | -35.5266 | -33.3096
22.2867 | -35.5271 | -33.316
22.2928 | -35.5277 | -33.3221
22.2989 | -35.5282 | -33.3287
22.305 | -35.5287 | -33.3345
22.3111 | -35.5293 | -33.3403
22.3172 | -35.5298 | -33.347
22.3233 | -35.5304 | -33.3537
22.3294 | -35.5309 | -33.36
22.3355 | -35.5314 | -33.3663
22.3416 | -35.5319 | -33.3727
22.3477 | -35.5325 | -33.3786
22.3538 | -35.533 | -33.3848
22.3599 | -35.5335 | -33.3913
22.366 | -35.534 | -33.3975
22.3721 | -35.5345 | -33.4038
22.3782 | -35.535 | -33.4102
22.3843 | -35.5355 | -33.4164
22.3904 | -35.536 | -33.4223
22.3965 | -35.5366 | -33.4289
22.4026 | -35.5371 | -33.4351
22.4087 | -35.5377 | -33.4416
22.4148 | -35.5382 | -33.4479
22.4209 | -35.5387 | -33.4542
22.4271 | -35.5393 | -33.4603
22.4332 | -35.5398 | -33.467
22.4393 | -35.5404 | -33.4732
22.4454 | -35.5409 | -33.4793
22.4515 | -35.5415 | -33.4858
22.4576 | -35.542 | -33.4922
22.4637 | -35.5426 | -33.4984
22.4698 | -35.5431 | -33.5045
22.4759 | -35.5436 | -33.5109
22.482 | -35.5441 | -33.5173
22.4881 | -35.5447 | -33.5236
22.4942 | -35.5452 | -33.5298
22.5003 | -35.5457 | -33.5363
22.5064 | -35.5463 | -33.5426
22.5125 | -35.5468 | -33.548
22.5186 | -35.5474 | -33.5541
22.5247 | -35.5478 | -33.5603
22.5308 | -35.5484 | -33.5663
22.5369 | -35.5488 | -33.5724
22.543 | -35.5494 | -33.5786
22.5491 | -35.55 | -33.5846
22.5552 | -35.5505 | -33.5907
22.5613 | -35.5511 | -33.5973
22.5675 | -35.5516 | -33.6034
22.5736 | -35.5522 | -33.6095
22.5797 | -35.5527 | -33.6158
22.5858 | -35.5532 | -33.6216
22.5919 | -35.5538 | -33.6278
22.598 | -35.5543 | -33.6343
22.6041 | -35.5549 | -33.6409
22.6102 | -35.5554 | -33.6471
22.6163 | -35.556 | -33.6535
22.6224 | -35.5566 | -33.6599
22.6285 | -35.5572 | -33.6662
22.6346 | -35.5577 | -33.6725
22.6407 | -35.5581 | -33.6785
22.6468 | -35.5586 | -33.6846
22.6529 | -35.5591 | -33.6904
22.659 | -35.5597 | -33.6969
22.6651 | -35.5602 | -33.7033
22.6712 | -35.5606 | -33.7094
22.6773 | -35.5612 | -33.7154
22.6834 | -35.5617 | -33.7219
22.6895 | -35.5622 | -33.7281
22.6956 | -35.5628 | -33.734
22.7017 | -35.5633 | -33.7407
22.7079 | -35.5637 | -33.7469
22.714 | -35.5643 | -33.7527
22.7201 | -35.5649 | -33.7587
22.7262 | -35.5654 | -33.7646
22.7323 | -35.566 | -33.7706
22.7384 | -35.5665 | -33.7768
22.7445 | -35.567 | -33.7831
22.7506 | -35.5675 | -33.7896
22.7567 | -35.5681 | -33.796
22.7628 | -35.5687 | -33.8022
22.7689 | -35.5693 | -33.8088
22.775 | -35.5699 | -33.8152
22.7811 | -35.5704 | -33.8212
22.7872 | -35.5709 | -33.8275
22.7933 | -35.5716 | -33.8335
22.7994 | -35.5721 | -33.8398
22.8055 | -35.5727 | -33.8466
22.8116 | -35.5732 | -33.853
22.8177 | -35.5737 | -33.8596
22.8238 | -35.5743 | -33.8654
22.8299 | -35.5748 | -33.8715
22.836 | -35.5754 | -33.8778
22.8421 | -35.5759 | -33.8841
22.8482 | -35.5764 | -33.8903
22.8544 | -35.577 | -33.8966
22.8605 | -35.5775 | -33.9025
22.8666 | -35.578 | -33.9084
22.8727 | -35.5785 | -33.9144
22.8788 | -35.5791 | -33.9203
22.8849 | -35.5796 | -33.9264
22.891 | -35.5801 | -33.9329
22.8971 | -35.5806 | -33.9393
22.9032 | -35.5812 | -33.9449
22.9093 | -35.5817 | -33.9515
22.9154 | -35.5822 | -33.9577
22.9215 | -35.5828 | -33.964
22.9276 | -35.5833 | -33.9702
22.9337 | -35.5838 | -33.9762
22.9398 | -35.5843 | -33.9823
22.9459 | -35.5849 | -33.9887
22.952 | -35.5853 | -33.9953
22.9581 | -35.5859 | -34.0018
22.9642 | -35.5864 | -34.0082
22.9703 | -35.587 | -34.0144
22.9764 | -35.5875 | -34.0207
22.9825 | -35.5881 | -34.027
22.9886 | -35.5886 | -34.0331
22.9948 | -35.5892 | -34.0393
23.0009 | -35.5896 | -34.0456
23.007 | -35.5902 | -34.0519
23.0131 | -35.5907 | -34.0585
23.0192 | -35.5912 | -34.0648
23.0253 | -35.5918 | -34.071
23.0314 | -35.5923 | -34.0771
23.0375 | -35.5929 | -34.0835
23.0436 | -35.5935 | -34.0902
23.0497 | -35.594 | -34.0964
23.0558 | -35.5946 | -34.1029
23.0619 | -35.5951 | -34.1095
23.068 | -35.5956 | -34.1157
23.0741 | -35.5961 | -34.1211
23.0802 | -35.5967 | -34.1275
23.0863 | -35.5972 | -34.1336
23.0924 | -35.5977 | -34.1398
23.0985 | -35.5982 | -34.1458
23.1046 | -35.5987 | -34.1524
23.1107 | -35.5992 | -34.159
23.1168 | -35.5998 | -34.1653
23.1229 | -35.6003 | -34.1717
23.129 | -35.6008 | -34.1781
23.1351 | -35.6014 | -34.1839
23.1413 | -35.6019 | -34.1904
23.1474 | -35.6023 | -34.1964
23.1535 | -35.6028 | -34.2021
23.1596 | -35.6034 | -34.2085
23.1657 | -35.6039 | -34.2143
23.1718 | -35.6044 | -34.2207
23.1779 | -35.605 | -34.2272
23.184 | -35.6055 | -34.2334
23.1901 | -35.606 | -34.2397
23.1962 | -35.6065 | -34.2458
23.2023 | -35.607 | -34.2521
23.2084 | -35.6075 | -34.2587
23.2145 | -35.608 | -34.2643
23.2206 | -35.6086 | -34.2706
23.2267 | -35.6091 | -34.2766
23.2328 | -35.6096 | -34.2826
23.2389 | -35.6101 | -34.2891
23.245 | -35.6107 | -34.2955
23.2511 | -35.6113 | -34.3021
23.2572 | -35.6118 | -34.3077
23.2633 | -35.6123 | -34.314
23.2694 | -35.6128 | -34.3202
23.2755 | -35.6133 | -34.3271
23.2817 | -35.6139 | -34.3333
23.2878 | -35.6144 | -34.3395
23.2939 | -35.6149 | -34.346
23.3 | -35.6154 | -34.3522
23.3061 | -35.6159 | -34.3582
23.3122 | -35.6165 | -34.3644
23.3183 | -35.617 | -34.3703
23.3244 | -35.6176 | -34.3765
23.3305 | -35.6181 | -34.3822
23.3366 | -35.6186 | -34.3882
23.3427 | -35.6191 | -34.3942
23.3488 | -35.6195 | -34.4003
23.3549 | -35.62 | -34.4067
23.361 | -35.6205 | -34.4129
23.3671 | -35.6211 | -34.4189
23.3732 | -35.6217 | -34.4248
23.3793 | -35.6222 | -34.4308
23.3854 | -35.6227 | -34.4374
23.3915 | -35.6233 | -34.4438
23.3976 | -35.6238 | -34.45
23.4037 | -35.6244 | -34.456
23.4098 | -35.6249 | -34.4618
23.4159 | -35.6255 | -34.4686
23.422 | -35.626 | -34.4749
23.4282 | -35.6265 | -34.4812
23.4343 | -35.6271 | -34.4873
23.4404 | -35.6276 | -34.4937
23.4465 | -35.6282 | -34.5
23.4526 | -35.6288 | -34.5062
23.4587 | -35.6293 | -34.5124
23.4648 | -35.6297 | -34.518
23.4709 | -35.6302 | -34.5243
23.477 | -35.6307 | -34.5303
23.4831 | -35.6312 | -34.5367
23.4892 | -35.6317 | -34.5431
23.4953 | -35.6323 | -34.5494
23.5014 | -35.6327 | -34.5557
23.5075 | -35.6332 | -34.5621
23.5136 | -35.6337 | -34.5682
23.5197 | -35.6342 | -34.5742
23.5258 | -35.6347 | -34.5802
23.5319 | -35.6352 | -34.5862
23.538 | -35.6357 | -34.5921
23.5441 | -35.6363 | -34.5984
23.5502 | -35.6367 | -34.6047
23.5563 | -35.6372 | -34.6106
23.5624 | -35.6377 | -34.6169
23.5686 | -35.6382 | -34.6225
23.5747 | -35.6388 | -34.629
23.5808 | -35.6393 | -34.6353
23.5869 | -35.6398 | -34.6414
23.593 | -35.6403 | -34.6474
23.5991 | -35.6408 | -34.6536
23.6052 | -35.6412 | -34.6599
23.6113 | -35.6417 | -34.6661
23.6174 | -35.6422 | -34.6724
23.6235 | -35.6428 | -34.6786
23.6296 | -35.6433 | -34.6847
23.6357 | -35.6439 | -34.6907
23.6418 | -35.6444 | -34.6968
23.6479 | -35.6449 | -34.7029
23.654 | -35.6455 | -34.7089
23.6601 | -35.6459 | -34.7145
23.6662 | -35.6464 | -34.7209
23.6723 | -35.6469 | -34.7265
23.6784 | -35.6474 | -34.7332
23.6845 | -35.648 | -34.7397
23.6906 | -35.6485 | -34.7462
23.6967 | -35.649 | -34.752
23.7028 | -35.6496 | -34.7586
23.7089 | -35.65 | -34.7644
23.7151 | -35.6506 | -34.7707
23.7212 | -35.651 | -34.7775
23.7273 | -35.6515 | -34.7835
23.7334 | -35.652 | -34.7891
23.7395 | -35.6525 | -34.7953
23.7456 | -35.6529 | -34.8012
23.7517 | -35.6535 | -34.8071
23.7578 | -35.6539 | -34.8125
23.7639 | -35.6544 | -34.8183
23.77 | -35.655 | -34.8246
23.7761 | -35.6555 | -34.8307
23.7822 | -35.656 | -34.8367
23.7883 | -35.6565 | -34.8436
23.7944 | -35.6571 | -34.8497
23.8005 | -35.6576 | -34.8558
23.8066 | -35.658 | -34.8619
23.8127 | -35.6586 | -34.8683
23.8188 | -35.6591 | -34.8746
23.8249 | -35.6596 | -34.8803
23.831 | -35.6601 | -34.887
23.8371 | -35.6605 | -34.893
23.8432 | -35.661 | -34.8995
23.8493 | -35.6615 | -34.9047
23.8555 | -35.6621 | -34.9113
23.8616 | -35.6625 | -34.9172
23.8677 | -35.663 | -34.9229
23.8738 | -35.6635 | -34.929
23.8799 | -35.6641 | -34.9353
23.886 | -35.6646 | -34.9407
23.8921 | -35.6651 | -34.9468
23.8982 | -35.6656 | -34.9528
23.9043 | -35.6661 | -34.9586
23.9104 | -35.6666 | -34.9644
23.9165 | -35.6671 | -34.9704
23.9226 | -35.6676 | -34.9759
23.9287 | -35.6681 | -34.9824
23.9348 | -35.6686 | -34.9884
23.9409 | -35.6692 | -34.9943
23.947 | -35.6698 | -35.0004
23.9531 | -35.6703 | -35.0059
23.9592 | -35.6708 | -35.0118
23.9653 | -35.6713 | -35.0186
23.9714 | -35.6718 | -35.0247
23.9775 | -35.6723 | -35.0306
23.9836 | -35.6728 | -35.0362
23.9897 | -35.6733 | -35.042
23.9958 | -35.6738 | -35.0481
24.002 | -35.6743 | -35.0544
24.0081 | -35.6749 | -35.0601
24.0142 | -35.6753 | -35.0658
24.0203 | -35.6758 | -35.0725
24.0264 | -35.6763 | -35.0775
24.0325 | -35.6768 | -35.0834
24.0386 | -35.6772 | -35.0894
24.0447 | -35.6777 | -35.0954
24.0508 | -35.6783 | -35.1015
24.0569 | -35.6787 | -35.1072
24.063 | -35.6792 | -35.1131
24.0691 | -35.6797 | -35.1189
24.0752 | -35.6802 | -35.125
24.0813 | -35.6807 | -35.1312
24.0874 | -35.6813 | -35.1372
24.0935 | -35.6818 | -35.1438
24.0996 | -35.6824 | -35.1493
24.1057 | -35.6829 | -35.1557
24.1118 | -35.6834 | -35.1617
24.1179 | -35.6839 | -35.1682
24.124 | -35.6845 | -35.1741
24.1301 | -35.685 | -35.1796
24.1362 | -35.6855 | -35.1853
24.1424 | -35.686 | -35.1915
24.1485 | -35.6865 | -35.1978
24.1546 | -35.687 | -35.2032
24.1607 | -35.6875 | -35.2088
24.1668 | -35.688 | -35.2151
24.1729 | -35.6885 | -35.2211
24.179 | -35.6891 | -35.2268
24.1851 | -35.6896 | -35.2331
24.1912 | -35.6901 | -35.2389
24.1973 | -35.6906 | -35.2447
24.2034 | -35.6912 | -35.2506
24.2095 | -35.6919 | -35.2562
24.2156 | -35.6924 | -35.263
24.2217 | -35.6928 | -35.269
24.2278 | -35.6934 | -35.2751
24.2339 | -35.694 | -35.2807
24.24 | -35.6944 | -35.2867
24.2461 | -35.6949 | -35.2928
24.2522 | -35.6954 | -35.2987
24.2583 | -35.6959 | -35.3049
24.2644 | -35.6964 | -35.3105
24.2705 | -35.6969 | -35.3163
24.2766 | -35.6974 | -35.3222
24.2827 | -35.698 | -35.3285
24.2889 | -35.6985 | -35.3342
24.295 | -35.6991 | -35.3399
24.3011 | -35.6995 | -35.3461
24.3072 | -35.7001 | -35.3513
24.3133 | -35.7006 | -35.3566
24.3194 | -35.7011 | -35.3623
24.3255 | -35.7018 | -35.3682
24.3316 | -35.7022 | -35.3747
24.3377 | -35.7027 | -35.3804
24.3438 | -35.7031 | -35.3858
24.3499 | -35.7037 | -35.3914
24.356 | -35.7041 | -35.3973
24.3621 | -35.7045 | -35.4035
24.3682 | -35.705 | -35.4093
24.3743 | -35.7054 | -35.4156
24.3804 | -35.7059 | -35.4214
24.3865 | -35.7064 | -35.4278
24.3926 | -35.7069 | -35.434
24.3987 | -35.7075 | -35.4407
24.4048 | -35.708 | -35.4465
24.4109 | -35.7085 | -35.452
24.417 | -35.709 | -35.4582
24.4231 | -35.7096 | -35.464
24.4293 | -35.7101 | -35.4699
24.4354 | -35.7106 | -35.4756
24.4415 | -35.7111 | -35.4816
24.4476 | -35.7118 | -35.4872
24.4537 | -35.7123 | -35.493
24.4598 | -35.7127 | -35.4995
24.4659 | -35.7132 | -35.5056
24.472 | -35.7137 | -35.5118
24.4781 | -35.7143 | -35.5178
24.4842 | -35.7146 | -35.5236
24.4903 | -35.7152 | -35.5295
24.4964 | -35.7157 | -35.5348
24.5025 | -35.7161 | -35.5407
24.5086 | -35.7167 | -35.5465
24.5147 | -35.7171 | -35.5519
24.5208 | -35.7176 | -35.5577
24.5269 | -35.7181 | -35.5639
24.533 | -35.7187 | -35.57
24.5391 | -35.7193 | -35.5759
24.5452 | -35.7197 | -35.5819
24.5513 | -35.7202 | -35.5875
24.5574 | -35.7207 | -35.5939
24.5635 | -35.7212 | -35.6002
24.5696 | -35.7216 | -35.6063
24.5758 | -35.7221 | -35.6124
24.5819 | -35.7226 | -35.6179
24.588 | -35.7231 | -35.6242
24.5941 | -35.7236 | -35.6302
24.6002 | -35.7241 | -35.6357
24.6063 | -35.7245 | -35.6414
24.6124 | -35.725 | -35.647
24.6185 | -35.7255 | -35.6525
24.6246 | -35.726 | -35.6583
24.6307 | -35.7265 | -35.6642
24.6368 | -35.7271 | -35.6698
24.6429 | -35.7275 | -35.6759
24.649 | -35.7281 | -35.6818
24.6551 | -35.7286 | -35.6877
24.6612 | -35.7291 | -35.6928
24.6673 | -35.7295 | -35.6988
24.6734 | -35.73 | -35.7047
24.6795 | -35.7305 | -35.7109
24.6856 | -35.7311 | -35.7168
24.6917 | -35.7316 | -35.7233
24.6978 | -35.732 | -35.728
24.7039 | -35.7326 | -35.734
24.71 | -35.7331 | -35.7397
24.7162 | -35.7335 | -35.7452
24.7223 | -35.7341 | -35.7512
24.7284 | -35.7345 | -35.7565
24.7345 | -35.735 | -35.7625
24.7406 | -35.7354 | -35.7685
24.7467 | -35.7359 | -35.7747
24.7528 | -35.7364 | -35.7805
24.7589 | -35.7369 | -35.7863
24.765 | -35.7373 | -35.792
24.7711 | -35.7378 | -35.7982
24.7772 | -35.7383 | -35.8036
24.7833 | -35.7389 | -35.809
24.7894 | -35.7394 | -35.8149
24.7955 | -35.7399 | -35.8203
24.8016 | -35.7405 | -35.8262
24.8077 | -35.741 | -35.8317
24.8138 | -35.7415 | -35.8377
24.8199 | -35.742 | -35.8434
24.826 | -35.7425 | -35.8487
24.8321 | -35.7431 | -35.8541
24.8382 | -35.7436 | -35.8597
24.8443 | -35.7442 | -35.8657
24.8504 | -35.7447 | -35.8709
24.8565 | -35.7453 | -35.877
24.8627 | -35.7457 | -35.8831
24.8688 | -35.7461 | -35.889
24.8749 | -35.7466 | -35.8946
24.881 | -35.7472 | -35.9005
24.8871 | -35.7477 | -35.9065
24.8932 | -35.7483 | -35.9124
24.8993 | -35.7488 | -35.9179
24.9054 | -35.7493 | -35.9236
24.9115 | -35.7499 | -35.9293
24.9176 | -35.7502 | -35.9353
24.9237 | -35.7507 | -35.9408
24.9298 | -35.7511 | -35.9462
24.9359 | -35.7516 | -35.9521
24.942 | -35.7521 | -35.9571
24.9481 | -35.7526 | -35.9636
24.9542 | -35.7531 | -35.9699
24.9603 | -35.7537 | -35.9764
24.9664 | -35.7542 | -35.9822
24.9725 | -35.7547 | -35.988
24.9786 | -35.7552 | -35.9931
24.9847 | -35.7557 | -35.9993
24.9908 | -35.7561 | -36.0053
24.9969 | -35.7567 | -36.0106
25.0031 | -35.7572 | -36.016
25.0092 | -35.7579 | -36.0214
25.0153 | -35.7583 | -36.0273
25.0214 | -35.7588 | -36.0326
25.0275 | -35.7593 | -36.0383
25.0336 | -35.76 | -36.0435
25.0397 | -35.7606 | -36.0485
25.0458 | -35.7612 | -36.0544
25.0519 | -35.7617 | -36.0598
25.058 | -35.7622 | -36.0654
25.0641 | -35.7627 | -36.0712
25.0702 | -35.7632 | -36.0772
25.0763 | -35.7638 | -36.0826
25.0824 | -35.7642 | -36.0884
25.0885 | -35.7647 | -36.0938
25.0946 | -35.7653 | -36.0996
25.1007 | -35.7657 | -36.1056
25.1068 | -35.7663 | -36.1113
25.1129 | -35.7667 | -36.1168
25.119 | -35.7673 | -36.1222
25.1251 | -35.7678 | -36.1276
25.1312 | -35.7684 | -36.1329
25.1373 | -35.7688 | -36.1384
25.1435 | -35.7693 | -36.1441
25.1496 | -35.7698 | -36.1494
25.1557 | -35.7702 | -36.1551
25.1618 | -35.7708 | -36.1608
25.1679 | -35.7713 | -36.1668
25.174 | -35.7718 | -36.1724
25.1801 | -35.7724 | -36.1782
25.1862 | -35.773 | -36.1835
25.1923 | -35.7734 | -36.1891
25.1984 | -35.7739 | -36.1946
25.2045 | -35.7744 | -36.2003
25.2106 | -35.7747 | -36.2061
25.2167 | -35.7752 | -36.2119
25.2228 | -35.7757 | -36.2175
25.2289 | -35.7762 | -36.2232
25.235 | -35.7767 | -36.2289
25.2411 | -35.7772 | -36.2346
25.2472 | -35.7776 | -36.24
25.2533 | -35.7781 | -36.2455
25.2594 | -35.7786 | -36.2505
25.2655 | -35.7792 | -36.2559
25.2716 | -35.7797 | -36.2617
25.2777 | -35.7802 | -36.2671
25.2838 | -35.7807 | -36.2729
25.29 | -35.7812 | -36.2785
25.2961 | -35.7817 | -36.2835
25.3022 | -35.7822 | -36.2886
25.3083 | -35.7827 | -36.2941
25.3144 | -35.7832 | -36.3
25.3205 | -35.7836 | -36.3057
25.3266 | -35.7841 | -36.3106
25.3327 | -35.7846 | -36.3162
25.3388 | -35.7851 | -36.3218
25.3449 | -35.7855 | -36.3274
25.351 | -35.786 | -36.333
25.3571 | -35.7865 | -36.3387
25.3632 | -35.7869 | -36.3436
25.3693 | -35.7873 | -36.3492
25.3754 | -35.7878 | -36.3547
25.3815 | -35.7882 | -36.3606
25.3876 | -35.7887 | -36.3661
25.3937 | -35.7892 | -36.3719
25.3998 | -35.7898 | -36.3768
25.4059 | -35.7903 | -36.3828
25.412 | -35.7907 | -36.3884
25.4181 | -35.7912 | -36.3941
25.4242 | -35.7917 | -36.3995
25.4304 | -35.7922 | -36.405
25.4365 | -35.7927 | -36.4102
25.4426 | -35.7932 | -36.4158
25.4487 | -35.7936 | -36.4218
25.4548 | -35.7941 | -36.4273
25.4609 | -35.7946 | -36.433
25.467 | -35.795 | -36.4386
25.4731 | -35.7956 | -36.4443
25.4792 | -35.7961 | -36.4498
25.4853 | -35.7966 | -36.4556
25.4914 | -35.7971 | -36.461
25.4975 | -35.7976 | -36.4666
25.5036 | -35.7981 | -36.4716
25.5097 | -35.7986 | -36.4771
25.5158 | -35.7991 | -36.4829
25.5219 | -35.7996 | -36.4885
25.528 | -35.8001 | -36.4943
25.5341 | -35.8006 | -36.5001
25.5402 | -35.8012 | -36.5052
25.5463 | -35.8017 | -36.5105
25.5524 | -35.8022 | -36.516
25.5585 | -35.8027 | -36.5217
25.5646 | -35.8032 | -36.5274
25.5707 | -35.8038 | -36.533
25.5769 | -35.8043 | -36.5384
25.583 | -35.8048 | -36.5438
25.5891 | -35.8054 | -36.5496
25.5952 | -35.8058 | -36.5549
25.6013 | -35.8063 | -36.5605
25.6074 | -35.8068 | -36.565
25.6135 | -35.8074 | -36.5702
25.6196 | -35.8079 | -36.576
25.6257 | -35.8083 | -36.5812
25.6318 | -35.8088 | -36.5866
25.6379 | -35.8093 | -36.5924
25.644 | -35.8098 | -36.5976
25.6501 | -35.8102 | -36.6033
25.6562 | -35.8107 | -36.6086
25.6623 | -35.8112 | -36.6138
25.6684 | -35.8117 | -36.6192
25.6745 | -35.8123 | -36.6247
25.6806 | -35.8127 | -36.6298
25.6867 | -35.8131 | -36.6356
25.6928 | -35.8136 | -36.641
25.6989 | -35.8141 | -36.6463
25.705 | -35.8147 | -36.6518
25.7111 | -35.8151 | -36.6569
25.7173 | -35.8156 | -36.6616
25.7234 | -35.8161 | -36.6671
25.7295 | -35.8166 | -36.6725
25.7356 | -35.817 | -36.6781
25.7417 | -35.8175 | -36.6837
25.7478 | -35.8179 | -36.6889
25.7539 | -35.8184 | -36.6943
25.76 | -35.8189 | -36.6997
25.7661 | -35.8195 | -36.7051
25.7722 | -35.82 | -36.7108
25.7783 | -35.8205 | -36.7163
25.7844 | -35.8209 | -36.7211
25.7905 | -35.8214 | -36.7268
25.7966 | -35.8219 | -36.7324
25.8027 | -35.8223 | -36.7377
25.8088 | -35.8228 | -36.7432
25.8149 | -35.8233 | -36.7485
25.821 | -35.8238 | -36.754
25.8271 | -35.8242 | -36.7595
25.8332 | -35.8247 | -36.7646
25.8393 | -35.8252 | -36.7707
25.8454 | -35.8257 | -36.7762
25.8515 | -35.8262 | -36.7816
25.8576 | -35.8267 | -36.7869
25.8638 | -35.8272 | -36.7924
25.8699 | -35.8276 | -36.798
25.876 | -35.8281 | -36.8034
25.8821 | -35.8286 | -36.8088
25.8882 | -35.8291 | -36.8141
25.8943 | -35.8296 | -36.8195
25.9004 | -35.83 | -36.8248
25.9065 | -35.8305 | -36.8305
25.9126 | -35.831 | -36.8358
25.9187 | -35.8315 | -36.8412
25.9248 | -35.8319 | -36.8462
25.9309 | -35.8324 | -36.8519
25.937 | -35.833 | -36.8578
25.9431 | -35.8334 | -36.8633
25.9492 | -35.834 | -36.8688
25.9553 | -35.8345 | -36.8745
25.9614 | -35.835 | -36.8792
25.9675 | -35.8356 | -36.8849
25.9736 | -35.8361 | -36.8906
25.9797 | -35.8366 | -36.8957
25.9858 | -35.837 | -36.901
25.9919 | -35.8375 | -36.9064
25.998 | -35.838 | -36.9115
26.0042 | -35.8385 | -36.917
26.0103 | -35.8389 | -36.9226
26.0164 | -35.8394 | -36.9281
26.0225 | -35.8399 | -36.9333
26.0286 | -35.8404 | -36.9386
26.0347 | -35.8409 | -36.9439
26.0408 | -35.8414 | -36.9494
26.0469 | -35.8419 | -36.955
26.053 | -35.8424 | -36.9605
26.0591 | -35.8429 | -36.9664
26.0652 | -35.8434 | -36.9717
26.0713 | -35.8438 | -36.9772
26.0774 | -35.8443 | -36.9824
26.0835 | -35.8448 | -36.9876
26.0896 | -35.8453 | -36.9928
26.0957 | -35.8458 | -36.9984
26.1018 | -35.8463 | -37.0037
26.1079 | -35.8467 | -37.009
26.114 | -35.8473 | -37.0146
26.1201 | -35.8477 | -37.02
26.1262 | -35.8482 | -37.0251
26.1323 | -35.8486 | -37.0304
26.1384 | -35.8491 | -37.0353
26.1445 | -35.8495 | -37.0406
26.1507 | -35.85 | -37.046
26.1568 | -35.8505 | -37.0513
26.1629 | -35.851 | -37.0566
26.169 | -35.8515 | -37.0617
26.1751 | -35.8519 | -37.0662
26.1812 | -35.8524 | -37.0716
26.1873 | -35.8529 | -37.0769
26.1934 | -35.8534 | -37.082
26.1995 | -35.854 | -37.0872
26.2056 | -35.8544 | -37.0923
26.2117 | -35.8549 | -37.0975
26.2178 | -35.8554 | -37.103
26.2239 | -35.8559 | -37.1082
26.23 | -35.8564 | -37.1139
26.2361 | -35.8569 | -37.1192
26.2422 | -35.8574 | -37.1246
26.2483 | -35.8579 | -37.1299
26.2544 | -35.8584 | -37.1351
26.2605 | -35.8588 | -37.1401
26.2666 | -35.8593 | -37.1454
26.2727 | -35.8598 | -37.1508
26.2788 | -35.8603 | -37.156
26.2849 | -35.8608 | -37.1617
26.2911 | -35.8613 | -37.1671
26.2972 | -35.8618 | -37.1723
26.3033 | -35.8622 | -37.1778
26.3094 | -35.8627 | -37.1828
26.3155 | -35.8632 | -37.188
26.3216 | -35.8637 | -37.1934
26.3277 | -35.8642 | -37.1987
26.3338 | -35.8647 | -37.2038
26.3399 | -35.8651 | -37.209
26.346 | -35.8656 | -37.214
26.3521 | -35.8661 | -37.2194
26.3582 | -35.8666 | -37.2244
26.3643 | -35.8671 | -37.2296
26.3704 | -35.8676 | -37.2349
26.3765 | -35.8681 | -37.2404
26.3826 | -35.8685 | -37.2453
26.3887 | -35.869 | -37.2504
26.3948 | -35.8696 | -37.2556
26.4009 | -35.8701 | -37.261
26.407 | -35.8706 | -37.2661
26.4131 | -35.8711 | -37.2714
26.4192 | -35.8715 | -37.2762
26.4253 | -35.872 | -37.2815
26.4314 | -35.8724 | -37.2867
26.4376 | -35.8729 | -37.2922
26.4437 | -35.8734 | -37.2973
26.4498 | -35.8739 | -37.3024
26.4559 | -35.8744 | -37.3074
26.462 | -35.8749 | -37.312
26.4681 | -35.8753 | -37.3169
26.4742 | -35.8758 | -37.3224
26.4803 | -35.8762 | -37.3277
26.4864 | -35.8767 | -37.3332
26.4925 | -35.8772 | -37.3384
26.4986 | -35.8777 | -37.3433
26.5047 | -35.8782 | -37.3486
26.5108 | -35.8787 | -37.3539
26.5169 | -35.8792 | -37.3589
26.523 | -35.8797 | -37.3639
26.5291 | -35.8802 | -37.3693
26.5352 | -35.8807 | -37.3745
26.5413 | -35.8812 | -37.38
26.5474 | -35.8817 | -37.3853
26.5535 | -35.8822 | -37.3906
26.5596 | -35.8826 | -37.3955
26.5657 | -35.8831 | -37.4007
26.5718 | -35.8836 | -37.4058
26.578 | -35.8841 | -37.4109
26.5841 | -35.8846 | -37.416
26.5902 | -35.885 | -37.4215
26.5963 | -35.8855 | -37.4259
26.6024 | -35.886 | -37.4311
26.6085 | -35.8865 | -37.4364
26.6146 | -35.887 | -37.4416
26.6207 | -35.8875 | -37.4468
26.6268 | -35.888 | -37.4522
26.6329 | -35.8884 | -37.457
26.639 | -35.8889 | -37.4622
26.6451 | -35.8895 | -37.4675
26.6512 | -35.89 | -37.473
26.6573 | -35.8906 | -37.4784
26.6634 | -35.891 | -37.4836
26.6695 | -35.8915 | -37.4889
26.6756 | -35.892 | -37.4941
26.6817 | -35.8925 | -37.4992
26.6878 | -35.8929 | -37.5043
26.6939 | -35.8934 | -37.5093
26.7 | -35.8939 | -37.5135
26.7061 | -35.8943 | -37.5186
26.7122 | -35.8949 | -37.5237
26.7183 | -35.8954 | -37.5291
26.7245 | -35.8958 | -37.5341
26.7306 | -35.8963 | -37.5391
26.7367 | -35.8968 | -37.5441
26.7428 | -35.8974 | -37.5492
26.7489 | -35.8979 | -37.5546
26.755 | -35.8983 | -37.5599
26.7611 | -35.8988 | -37.5653
26.7672 | -35.8993 | -37.5706
26.7733 | -35.8997 | -37.5756
26.7794 | -35.9002 | -37.5809
26.7855 | -35.9007 | -37.5861
26.7916 | -35.9012 | -37.5915
26.7977 | -35.9017 | -37.5966
26.8038 | -35.9022 | -37.6018
26.8099 | -35.9027 | -37.6074
26.816 | -35.9032 | -37.6125
26.8221 | -35.9037 | -37.6179
26.8282 | -35.9042 | -37.6231
26.8343 | -35.9047 | -37.6282
26.8404 | -35.9052 | -37.6329
26.8465 | -35.9057 | -37.638
26.8526 | -35.9061 | -37.643
26.8587 | -35.9066 | -37.6476
26.8649 | -35.9071 | -37.6526
26.871 | -35.9076 | -37.6581
26.8771 | -35.908 | -37.6629
26.8832 | -35.9084 | -37.6681
26.8893 | -35.9088 | -37.6735
26.8954 | -35.9093 | -37.6788
26.9015 | -35.9099 | -37.6835
26.9076 | -35.9104 | -37.6888
26.9137 | -35.9109 | -37.6938
26.9198 | -35.9114 | -37.6988
26.9259 | -35.9118 | -37.7037
26.932 | -35.9123 | -37.7089
26.9381 | -35.9128 | -37.7138
26.9442 | -35.9132 | -37.7191
26.9503 | -35.9136 | -37.7242
26.9564 | -35.9141 | -37.7297
26.9625 | -35.9146 | -37.7347
26.9686 | -35.915 | -37.7398
26.9747 | -35.9155 | -37.7452
26.9808 | -35.9159 | -37.7497
26.9869 | -35.9163 | -37.755
26.993 | -35.9168 | -37.76
26.9991 | -35.9173 | -37.765
27.0052 | -35.9177 | -37.7698
27.0114 | -35.9182 | -37.7749
27.0175 | -35.9185 | -37.7796
27.0236 | -35.919 | -37.7847
27.0297 | -35.9195 | -37.7895
27.0358 | -35.9199 | -37.7947
27.0419 | -35.9204 | -37.7996
27.048 | -35.9208 | -37.8051
27.0541 | -35.9213 | -37.81
27.0602 | -35.9218 | -37.8149
27.0663 | -35.9223 | -37.8202
27.0724 | -35.9228 | -37.8251
27.0785 | -35.9233 | -37.8304
27.0846 | -35.9237 | -37.8352
27.0907 | -35.9242 | -37.84
27.0968 | -35.9246 | -37.8452
27.1029 | -35.925 | -37.8505
27.109 | -35.9255 | -37.8556
27.1151 | -35.926 | -37.8608
27.1212 | -35.9264 | -37.8656
27.1273 | -35.9269 | -37.8707
27.1334 | -35.9274 | -37.876
27.1395 | -35.9278 | -37.8814
27.1456 | -35.9283 | -37.8866
27.1518 | -35.9288 | -37.8916
27.1579 | -35.9293 | -37.8964
27.164 | -35.9298 | -37.9014
27.1701 | -35.9303 | -37.9065
27.1762 | -35.9308 | -37.9115
27.1823 | -35.9312 | -37.9168
27.1884 | -35.9316 | -37.922
27.1945 | -35.9322 | -37.9269
27.2006 | -35.9326 | -37.9319
27.2067 | -35.933 | -37.9367
27.2128 | -35.9334 | -37.9418
27.2189 | -35.9339 | -37.9472
27.225 | -35.9344 | -37.9523
27.2311 | -35.9349 | -37.9573
27.2372 | -35.9354 | -37.9624
27.2433 | -35.9359 | -37.9671
27.2494 | -35.9364 | -37.9721
27.2555 | -35.9369 | -37.9772
27.2616 | -35.9373 | -37.9821
27.2677 | -35.9377 | -37.9869
27.2738 | -35.9382 | -37.9917
27.2799 | -35.9387 | -37.9969
27.286 | -35.9391 | -38.0019
27.2921 | -35.9395 | -38.0074
27.2983 | -35.94 | -38.0122
27.3044 | -35.9405 | -38.0174
27.3105 | -35.9409 | -38.0226
27.3166 | -35.9414 | -38.0278
27.3227 | -35.9418 | -38.0328
27.3288 | -35.9423 | -38.0384
27.3349 | -35.9428 | -38.0434
27.341 | -35.9432 | -38.0486
27.3471 | -35.9437 | -38.0536
27.3532 | -35.9441 | -38.0585
27.3593 | -35.9446 | -38.0636
27.3654 | -35.945 | -38.0687
27.3715 | -35.9455 | -38.0732
27.3776 | -35.946 | -38.0782
27.3837 | -35.9465 | -38.0833
27.3898 | -35.9469 | -38.0883
27.3959 | -35.9474 | -38.0934
27.402 | -35.9478 | -38.0984
27.4081 | -35.9483 | -38.1034
27.4142 | -35.9488 | -38.1087
27.4203 | -35.9492 | -38.1138
27.4264 | -35.9497 | -38.1188
27.4325 | -35.9501 | -38.1239
27.4387 | -35.9506 | -38.1288
27.4448 | -35.951 | -38.1337
27.4509 | -35.9514 | -38.1387
27.457 | -35.9519 | -38.1438
27.4631 | -35.9523 | -38.1489
27.4692 | -35.9528 | -38.1543
27.4753 | -35.9532 | -38.1593
27.4814 | -35.9536 | -38.1644
27.4875 | -35.954 | -38.1694
27.4936 | -35.9545 | -38.1744
27.4997 | -35.955 | -38.1793
27.5058 | -35.9555 | -38.1846
27.5119 | -35.9559 | -38.1896
27.518 | -35.9563 | -38.1944
27.5241 | -35.9567 | -38.1994
27.5302 | -35.9571 | -38.2048
27.5363 | -35.9576 | -38.2097
27.5424 | -35.958 | -38.2148
27.5485 | -35.9584 | -38.2201
27.5546 | -35.9589 | -38.2251
27.5607 | -35.9594 | -38.2302
27.5668 | -35.9598 | -38.2351
27.5729 | -35.9603 | -38.2401
27.5791 | -35.9608 | -38.2449
27.5852 | -35.9612 | -38.2501
27.5913 | -35.9616 | -38.255
27.5974 | -35.9621 | -38.26
27.6035 | -35.9626 | -38.265
27.6096 | -35.963 | -38.2702
27.6157 | -35.9634 | -38.2752
27.6218 | -35.9639 | -38.2803
27.6279 | -35.9644 | -38.2853
27.634 | -35.9647 | -38.2904
27.6401 | -35.9652 | -38.2955
27.6462 | -35.9657 | -38.3004
27.6523 | -35.9661 | -38.3048
27.6584 | -35.9665 | -38.31
27.6645 | -35.967 | -38.3149
27.6706 | -35.9674 | -38.3198
27.6767 | -35.9678 | -38.3247
27.6828 | -35.9682 | -38.3298
27.6889 | -35.9686 | -38.3348
27.695 | -35.9691 | -38.34
27.7011 | -35.9695 | -38.3445
27.7072 | -35.97 | -38.3494
27.7133 | -35.9705 | -38.3547
27.7194 | -35.9709 | -38.3597
27.7256 | -35.9714 | -38.3651
27.7317 | -35.9718 | -38.3699
27.7378 | -35.9722 | -38.3749
27.7439 | -35.9727 | -38.3803
27.75 | -35.9732 | -38.3851
27.7561 | -35.9735 | -38.3901
27.7622 | -35.974 | -38.3952
27.7683 | -35.9745 | -38.4001
27.7744 | -35.9749 | -38.4046
27.7805 | -35.9753 | -38.4099
27.7866 | -35.9758 | -38.4152
27.7927 | -35.9762 | -38.4203
27.7988 | -35.9767 | -38.4253
27.8049 | -35.9772 | -38.4308
27.811 | -35.9776 | -38.4359
27.8171 | -35.9781 | -38.4406
27.8232 | -35.9784 | -38.4455
27.8293 | -35.9789 | -38.4506
27.8354 | -35.9794 | -38.4559
27.8415 | -35.9798 | -38.4609
27.8476 | -35.9802 | -38.4663
27.8537 | -35.9806 | -38.4714
27.8598 | -35.981 | -38.4761
27.866 | -35.9814 | -38.4814
27.8721 | -35.9818 | -38.4863
27.8782 | -35.9822 | -38.4916
27.8843 | -35.9827 | -38.4966
27.8904 | -35.9831 | -38.5014
27.8965 | -35.9836 | -38.5063
27.9026 | -35.984 | -38.5113
27.9087 | -35.9845 | -38.5164
27.9148 | -35.9849 | -38.5213
27.9209 | -35.9853 | -38.5261
27.927 | -35.9857 | -38.5312
27.9331 | -35.9861 | -38.5362
27.9392 | -35.9865 | -38.5416
27.9453 | -35.987 | -38.5463
27.9514 | -35.9874 | -38.5517
27.9575 | -35.9877 | -38.557
27.9636 | -35.9881 | -38.5622
27.9697 | -35.9885 | -38.5672
27.9758 | -35.9889 | -38.5727
27.9819 | -35.9894 | -38.5777
27.988 | -35.9898 | -38.5827
27.9941 | -35.9902 | -38.5878
28.0002 | -35.9907 | -38.593
28.0063 | -35.9911 | -38.598
28.0125 | -35.9916 | -38.603
28.0186 | -35.992 | -38.6081
28.0247 | -35.9924 | -38.6132
28.0308 | -35.9929 | -38.6179
28.0369 | -35.9933 | -38.6228
28.043 | -35.9937 | -38.6278
28.0491 | -35.9942 | -38.6327
28.0552 | -35.9946 | -38.638
28.0613 | -35.995 | -38.6429
28.0674 | -35.9954 | -38.648
28.0735 | -35.9959 | -38.6531

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.6948 | -42.9576
21.4908 | -35.6299 | -35.7658
16.4948 | -34.9617 | -28.2566
12.6602 | -34.5706 | -21.8649
9.7171 | -34.3334 | -16.8749
7.4582 | -34.1752 | -12.8186
5.7244 | -34.0613 | -9.5349
4.3936 | -33.9836 | -6.8858
3.3722 | -33.9697 | -4.7986
2.5883 | -33.9658 | -3.2118
1.9866 | -33.9583 | -1.9467
1.5248 | -33.9589 | -0.93328
1.1703 | -33.9651 | -0.26812
1 | -33.9637 | 0.063012
0.89822 | -33.9616 | 0.33883
0.68941 | -33.9401 | 0.93139
0.52919 | -33.9449 | 1.3877
0.40612 | -33.9353 | 2.0028
0.31171 | -33.9357 | 2.462
0.23928 | -33.9275 | 3.1873
0.18358 | -33.9443 | 3.6043
0.14091 | -33.9686 | 4.481
0.10819 | -33.9934 | 5.6404
0.083017 | -34.0007 | 7.1286
0.063705 | -34.0537 | 9.001
0.048923 | -34.117 | 11.508
0.037479 | -34.2196 | 14.711
0.028801 | -34.4054 | 18.7409
0.022078 | -34.7152 | 24.1204
0.017023 | -35.1948 | 30.5285
0.013018 | -36.0082 | 38.2066
0.010014 | -37.1185 | 46.1193

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.00378 | -1.7622
0.007559 | -29.7994
0.011339 | -35.3184
0.015119 | -41.3738
0.018898 | -44.3369
0.022678 | -46.5767
0.026458 | -48.2118
0.030237 | -49.1215
0.034017 | -50.6248
0.037796 | -51.7988
0.041576 | -52.7426
0.045356 | -52.5029
0.049135 | -53.2593
0.052915 | -55.5313
0.056695 | -57.0214
0.060474 | -58.1617
0.064254 | -60.0463
0.068034 | -60.1202
0.071813 | -59.4416
0.075593 | -59.7777
0.079373 | -60.8543
0.083152 | -61.599
0.086932 | -61.292
0.090712 | -62.0174
0.094491 | -62.7479
0.098271 | -63.9717
0.10205 | -64.2337
0.10583 | -63.2315
0.10961 | -63.262
0.11339 | -64.3416
0.11717 | -64.3927
0.12095 | -65.834
0.12473 | -66.5139
0.12851 | -66.6642
0.13229 | -66.006
0.13607 | -66.042
0.13985 | -66.4511
0.14363 | -66.886
0.14741 | -67.8752
0.15119 | -67.9819
0.15497 | -67.6942
0.15874 | -67.874
0.16253 | -69.0461
0.1663 | -69.4787
0.17008 | -69.1009
0.17386 | -69.2755
0.17764 | -69.8477
0.18142 | -70.0049
0.1852 | -70.3655
0.18898 | -70.4557
0.19276 | -71.4102
0.19654 | -71.4495
0.20032 | -70.155
0.2041 | -70.2222
0.20788 | -70.8995
0.21166 | -71.6489
0.21544 | -71.6392
0.21922 | -71.5876
0.223 | -72.0087
0.22678 | -71.6848
0.23056 | -72.0172
0.23434 | -72.4017
0.23812 | -71.7679
0.2419 | -71.0631
0.24568 | -72.2035
0.24946 | -72.9135
0.25324 | -73.6007
0.25702 | -73.3646
0.2608 | -72.5009
0.26458 | -72.7197
0.26836 | -73.0412
0.27214 | -74.4737
0.27591 | -73.9629
0.27969 | -74.7192
0.28347 | -74.9771
0.28725 | -74.8154
0.29103 | -74.9585
0.29481 | -74.6042
0.29859 | -74.8829
0.30237 | -74.7978
0.30615 | -75.4215
0.30993 | -76.5687
0.31371 | -76.2658
0.31749 | -75.885
0.32127 | -76.3133
0.32505 | -77.2759
0.32883 | -76.5105
0.33261 | -75.8097
0.33639 | -75.5461
0.34017 | -75.4555
0.34395 | -76.1932
0.34773 | -76.9781
0.35151 | -77.3379
0.35529 | -77.0196
0.35907 | -77.0808
0.36285 | -76.4485
0.36663 | -76.6636
0.3704 | -76.3841
0.37418 | -76.4722
0.37796 | -77.6294
0.38174 | -77.8741
0.38552 | -78.1623
0.3893 | -78.4425
0.39308 | -77.7229
0.39686 | -78.1112
0.40064 | -79.2089
0.40442 | -79.106
0.4082 | -78.9391
0.41198 | -78.1362
0.41576 | -77.6515
0.41954 | -77.6982
0.42332 | -79.1238
0.4271 | -78.5441
0.43088 | -78.6575
0.43466 | -79.4697
0.43844 | -79.2037
0.44222 | -77.9676
0.446 | -78.2943
0.44978 | -78.9332
0.45356 | -79.6728
0.45734 | -79.2183
0.46112 | -78.6543
0.4649 | -78.8776
0.46868 | -78.205
0.47246 | -77.7998
0.47624 | -79.0722
0.48002 | -79.8028
0.48379 | -79.0981
0.48757 | -79.1731
0.49135 | -78.5173
0.49513 | -78.6992
0.49891 | -78.82
0.50269 | -78.8637
0.50647 | -79.592
0.51025 | -79.4158
0.51403 | -79.9709
0.51781 | -79.8625
0.52159 | -79.8731
0.52537 | -80.8609
0.52915 | -80.2666
0.53293 | -80.142
0.53671 | -80.3782
0.54049 | -80.012
0.54427 | -80.3389
0.54805 | -80.7867
0.55183 | -81.1116
0.55561 | -80.964
0.55939 | -80.7027
0.56317 | -80.6456
0.56695 | -80.5931
0.57073 | -80.5935
0.57451 | -80.8854
0.57829 | -81.261
0.58207 | -81.4463
0.58584 | -81.1236
0.58962 | -80.613
0.5934 | -80.6565
0.59718 | -81.1465
0.60096 | -81.3587
0.60474 | -81.5094
0.60852 | -81.718
0.6123 | -81.5526
0.61608 | -81.5525
0.61986 | -81.6856
0.62364 | -81.804
0.62742 | -81.9951
0.6312 | -81.7754
0.63498 | -81.2311
0.63876 | -80.9071
0.64254 | -80.7876
0.64632 | -80.9875
0.6501 | -81.2665
0.65388 | -81.2538
0.65766 | -81.6251
0.66144 | -82.3425
0.66522 | -82.901
0.669 | -82.786
0.67278 | -82.4876
0.67656 | -82.7821
0.68034 | -83.0927
0.68412 | -82.9275
0.6879 | -82.8991
0.69168 | -83.086
0.69546 | -82.9821
0.69924 | -82.5052
0.70301 | -82.4462
0.70679 | -82.87
0.71057 | -83.0008
0.71435 | -82.7435
0.71813 | -82.2875
0.72191 | -82.2458
0.72569 | -82.4764
0.72947 | -82.6554
0.73325 | -83.191
0.73703 | -83.8834
0.74081 | -84.1289
0.74459 | -84.1384
0.74837 | -84.2308
0.75215 | -84.3051
0.75593 | -84.3695
0.75971 | -84.6054
0.76349 | -84.5284
0.76727 | -84.0234
0.77105 | -83.4159
0.77483 | -83.2195
0.77861 | -83.3802
0.78239 | -83.2371
0.78617 | -83.0678
0.78995 | -83.0627
0.79373 | -82.9668
0.79751 | -83.1748
0.80129 | -83.9259
0.80507 | -84.7581
0.80884 | -85.1837
0.81262 | -85.1607
0.8164 | -84.9659
0.82018 | -84.6049
0.82396 | -84.1846
0.82774 | -84.1392
0.83152 | -84.1856
0.8353 | -84.0561
0.83908 | -84.29
0.84286 | -84.7025
0.84664 | -84.587
0.85042 | -84.3249
0.8542 | -84.1071
0.85798 | -83.9676
0.86176 | -84.0352
0.86554 | -84.3243
0.86932 | -84.5426
0.8731 | -84.9477
0.87688 | -85.6443
0.88066 | -85.8222
0.88444 | -85.2639
0.88822 | -84.8257
0.892 | -84.8136
0.89578 | -85.081
0.89956 | -85.5075
0.90334 | -85.7401
0.90712 | -85.514
0.9109 | -85.1399
0.91467 | -84.8164
0.91845 | -84.888
0.92223 | -85.6039
0.92601 | -86.0332
0.92979 | -85.8989
0.93357 | -85.2735
0.93735 | -84.6678
0.94113 | -84.9202
0.94491 | -85.7491
0.94869 | -85.984
0.95247 | -85.812
0.95625 | -85.7607
0.96003 | -85.8941
0.96381 | -86.3498
0.96759 | -86.7311
0.97137 | -86.5146
0.97515 | -85.6834
0.97893 | -85.0307
0.98271 | -85.2847
0.98649 | -85.8551
0.99027 | -85.8395
0.99405 | -85.2915
0.99783 | -85.068
1.0016 | -85.091
1.0054 | -85.1972
1.0092 | -85.2322
1.0129 | -85.1775
1.0167 | -85.2835
1.0205 | -85.7804
1.0243 | -86.518
1.0281 | -86.6605
1.0318 | -86.6071
1.0356 | -86.46
1.0394 | -86.1765
1.0432 | -86.0471
1.047 | -85.8873
1.0507 | -86.0386
1.0545 | -86.4939
1.0583 | -86.5693
1.0621 | -86.0634
1.0659 | -85.6416
1.0696 | -85.96
1.0734 | -87.0583
1.0772 | -87.7898
1.081 | -87.6872
1.0848 | -87.2849
1.0885 | -87.0627
1.0923 | -86.9182
1.0961 | -86.9357
1.0999 | -86.9822
1.1037 | -86.9034
1.1074 | -86.8194
1.1112 | -86.9592
1.115 | -87.3535
1.1188 | -87.7063
1.1226 | -87.7728
1.1263 | -87.7139
1.1301 | -87.6451
1.1339 | -87.5536
1.1377 | -87.5125
1.1415 | -87.5292
1.1452 | -87.4824
1.149 | -87.4471
1.1528 | -87.6255
1.1566 | -87.9646
1.1604 | -88.1623
1.1641 | -88.0154
1.1679 | -87.6678
1.1717 | -87.443
1.1755 | -87.458
1.1792 | -87.5627
1.183 | -87.5831
1.1868 | -87.6027
1.1906 | -87.8663
1.1944 | -88.3801
1.1981 | -88.8674
1.2019 | -89.1287
1.2057 | -89.0595
1.2095 | -88.6012
1.2133 | -87.9934
1.217 | -87.5661
1.2208 | -87.4158
1.2246 | -87.4795
1.2284 | -87.7235
1.2322 | -88.0757
1.2359 | -88.4018
1.2397 | -88.6504
1.2435 | -88.7994
1.2473 | -88.86
1.2511 | -88.8634
1.2548 | -88.7076
1.2586 | -88.4263
1.2624 | -88.2859
1.2662 | -88.4784
1.27 | -88.8077
1.2737 | -88.8157
1.2775 | -88.4552
1.2813 | -88.1817
1.2851 | -88.307
1.2889 | -88.7201
1.2926 | -89.028
1.2964 | -89.0417
1.3002 | -88.8806
1.304 | -88.7087
1.3078 | -88.6568
1.3115 | -88.715
1.3153 | -88.7154
1.3191 | -88.6456
1.3229 | -88.6449
1.3267 | -88.7394
1.3304 | -88.8366
1.3342 | -88.9944
1.338 | -89.4045
1.3418 | -89.9757
1.3456 | -90.2787
1.3493 | -90.0507
1.3531 | -89.5685
1.3569 | -89.2659
1.3607 | -89.2773
1.3645 | -89.4869
1.3682 | -89.702
1.372 | -89.8395
1.3758 | -89.9225
1.3796 | -89.8178
1.3834 | -89.5051
1.3871 | -89.2124
1.3909 | -89.1598
1.3947 | -89.3981
1.3985 | -89.6838
1.4022 | -89.6521
1.406 | -89.356
1.4098 | -89.0991
1.4136 | -89.0481
1.4174 | -89.1802
1.4211 | -89.4015
1.4249 | -89.7249
1.4287 | -90.1256
1.4325 | -90.3532
1.4363 | -90.2422
1.44 | -89.8918
1.4438 | -89.6167
1.4476 | -89.6497
1.4514 | -89.884
1.4552 | -90.0695
1.4589 | -90.0687
1.4627 | -89.8851
1.4665 | -89.6689
1.4703 | -89.6052
1.4741 | -89.8332
1.4778 | -90.3705
1.4816 | -90.8937
1.4854 | -90.9609
1.4892 | -90.6351
1.493 | -90.2579
1.4967 | -89.9785
1.5005 | -89.7485
1.5043 | -89.5717
1.5081 | -89.5409
1.5119 | -89.6053
1.5156 | -89.6798
1.5194 | -89.8702
1.5232 | -90.236
1.527 | -90.5943
1.5308 | -90.7918
1.5345 | -90.8448
1.5383 | -90.6994
1.5421 | -90.4392
1.5459 | -90.2489
1.5497 | -90.1917
1.5534 | -90.3066
1.5572 | -90.53
1.561 | -90.6726
1.5648 | -90.6617
1.5686 | -90.5688
1.5723 | -90.5153
1.5761 | -90.52
1.5799 | -90.5419
1.5837 | -90.6537
1.5875 | -90.9307
1.5912 | -91.168
1.595 | -91.0993
1.5988 | -90.9278
1.6026 | -90.8735
1.6063 | -90.8049
1.6101 | -90.6974
1.6139 | -90.666
1.6177 | -90.7336
1.6215 | -90.8832
1.6252 | -91.043
1.629 | -91.0652
1.6328 | -90.9099
1.6366 | -90.8328
1.6404 | -90.8031
1.6441 | -90.8249
1.6479 | -90.8998
1.6517 | -91.0689
1.6555 | -91.3276
1.6593 | -91.6113
1.663 | -91.8336
1.6668 | -91.9371
1.6706 | -91.9647
1.6744 | -92.0028
1.6782 | -92.0862
1.6819 | -92.1862
1.6857 | -92.2354
1.6895 | -92.2262
1.6933 | -92.1661
1.6971 | -92.0653
1.7008 | -91.9062
1.7046 | -91.6996
1.7084 | -91.5362
1.7122 | -91.4745
1.716 | -91.5085
1.7197 | -91.5658
1.7235 | -91.5957
1.7273 | -91.5944
1.7311 | -91.5777
1.7349 | -91.5522
1.7386 | -91.4902
1.7424 | -91.3782
1.7462 | -91.2862
1.75 | -91.2883
1.7538 | -91.3886
1.7575 | -91.4929
1.7613 | -91.5271
1.7651 | -91.4799
1.7689 | -91.3975
1.7727 | -91.3615
1.7764 | -91.3896
1.7802 | -91.4423
1.784 | -91.5322
1.7878 | -91.7069
1.7916 | -91.9588
1.7953 | -92.2005
1.7991 | -92.3621
1.8029 | -92.4536
1.8067 | -92.4892
1.8105 | -92.435
1.8142 | -92.2768
1.818 | -92.0766
1.8218 | -91.9453
1.8256 | -91.928
1.8293 | -91.9772
1.8331 | -91.9956
1.8369 | -91.9348
1.8407 | -91.8568
1.8445 | -91.7824
1.8482 | -91.6812
1.852 | -91.6112
1.8558 | -91.6969
1.8596 | -91.9481
1.8634 | -92.2036
1.8671 | -92.3336
1.8709 | -92.3688
1.8747 | -92.4141
1.8785 | -92.492
1.8823 | -92.5221
1.886 | -92.4359
1.8898 | -92.2468
1.8936 | -92.0664
1.8974 | -92.0012
1.9012 | -92.0534
1.9049 | -92.1794
1.9087 | -92.3227
1.9125 | -92.4647
1.9163 | -92.5129
1.9201 | -92.4358
1.9238 | -92.323
1.9276 | -92.2917
1.9314 | -92.4149
1.9352 | -92.6429
1.939 | -92.8271
1.9427 | -92.8567
1.9465 | -92.7341
1.9503 | -92.5494
1.9541 | -92.3587
1.9579 | -92.2137
1.9616 | -92.1572
1.9654 | -92.199
1.9692 | -92.2853
1.973 | -92.3294
1.9768 | -92.3236
1.9805 | -92.3258
1.9843 | -92.3603
1.9881 | -92.4038
1.9919 | -92.4394
1.9957 | -92.4556
1.9994 | -92.4635
2.0032 | -92.4988
2.007 | -92.5595
2.0108 | -92.6171
2.0146 | -92.6701
2.0183 | -92.7169
2.0221 | -92.7399
2.0259 | -92.7365
2.0297 | -92.7167
2.0335 | -92.6886
2.0372 | -92.6976
2.041 | -92.7853
2.0448 | -92.9741
2.0486 | -93.2391
2.0523 | -93.4762
2.0561 | -93.5399
2.0599 | -93.3637
2.0637 | -93.0777
2.0675 | -92.8537
2.0712 | -92.7681
2.075 | -92.7874
2.0788 | -92.8439
2.0826 | -92.8986
2.0864 | -92.9256
2.0901 | -92.9419
2.0939 | -93.0066
2.0977 | -93.1229
2.1015 | -93.2418
2.1053 | -93.3211
2.109 | -93.3584
2.1128 | -93.3568
2.1166 | -93.3832
2.1204 | -93.5126
2.1242 | -93.6984
2.1279 | -93.8246
2.1317 | -93.8189
2.1355 | -93.7096
2.1393 | -93.5522
2.1431 | -93.4005
2.1468 | -93.3166
2.1506 | -93.3028
2.1544 | -93.3018
2.1582 | -93.2927
2.162 | -93.3023
2.1657 | -93.3481
2.1695 | -93.3888
2.1733 | -93.391
2.1771 | -93.3869
2.1809 | -93.4902
2.1846 | -93.5588
2.1884 | -93.5924
2.1922 | -93.5906
2.196 | -93.5782
2.1998 | -93.594
2.2035 | -93.6397
2.2073 | -93.6824
2.2111 | -93.6988
2.2149 | -93.6762
2.2187 | -93.6099
2.2224 | -93.5474
2.2262 | -93.5152
2.23 | -93.4983
2.2338 | -93.4592
2.2376 | -93.3903
2.2413 | -93.2994
2.2451 | -93.1972
2.2489 | -93.1179
2.2527 | -93.1046
2.2564 | -93.1718
2.2602 | -93.2774
2.264 | -93.3774
2.2678 | -93.467
2.2716 | -93.5565
2.2753 | -93.6642
2.2791 | -93.765
2.2829 | -93.8004
2.2867 | -93.7399
2.2905 | -93.5708
2.2942 | -93.347
2.298 | -93.1685
2.3018 | -93.1021
2.3056 | -93.1542
2.3094 | -93.2821
2.3131 | -93.4374
2.3169 | -93.547
2.3207 | -93.5867
2.3245 | -93.5775
2.3283 | -93.557
2.332 | -93.5781
2.3358 | -93.6356
2.3396 | -93.7002
2.3434 | -93.7627
2.3472 | -93.8388
2.3509 | -93.9417
2.3547 | -94.0718
2.3585 | -94.2271
2.3623 | -94.3735
2.3661 | -94.4804
2.3698 | -94.5385
2.3736 | -94.5452
2.3774 | -94.5055
2.3812 | -94.4345
2.385 | -94.3715
2.3887 | -94.3588
2.3925 | -94.3949
2.3963 | -94.4639
2.4001 | -94.5072
2.4039 | -94.5001
2.4076 | -94.4681
2.4114 | -94.4271
2.4152 | -94.3911
2.419 | -94.3675
2.4228 | -94.3506
2.4265 | -94.3341
2.4303 | -94.3062
2.4341 | -94.2599
2.4379 | -94.1823
2.4417 | -94.0858
2.4454 | -93.9751
2.4492 | -93.8521
2.453 | -93.7474
2.4568 | -93.6989
2.4606 | -93.7427
2.4643 | -93.8834
2.4681 | -94.0965
2.4719 | -94.3591
2.4757 | -94.6319
2.4794 | -94.8578
2.4832 | -94.983
2.487 | -94.9816
2.4908 | -94.8994
2.4946 | -94.8194
2.4983 | -94.7735
2.5021 | -94.7392
2.5059 | -94.6864
2.5097 | -94.6119
2.5135 | -94.5148
2.5172 | -94.4228
2.521 | -94.3618
2.5248 | -94.3346
2.5286 | -94.3409
2.5324 | -94.3707
2.5361 | -94.3958
2.5399 | -94.4148
2.5437 | -94.4363
2.5475 | -94.4736
2.5513 | -94.5283
2.555 | -94.6235
2.5588 | -94.7599
2.5626 | -94.8998
2.5664 | -94.9915
2.5702 | -94.9831
2.5739 | -94.8698
2.5777 | -94.6939
2.5815 | -94.4791
2.5853 | -94.2484
2.5891 | -94.0282
2.5928 | -93.8446
2.5966 | -93.7206
2.6004 | -93.6727
2.6042 | -93.7209
2.608 | -93.8635
2.6117 | -94.0818
2.6155 | -94.3409
2.6193 | -94.5717
2.6231 | -94.7487
2.6269 | -94.8662
2.6306 | -94.907
2.6344 | -94.8621
2.6382 | -94.7389
2.642 | -94.5876
2.6458 | -94.4674
2.6495 | -94.4417
2.6533 | -94.5458
2.6571 | -94.7626
2.6609 | -95.0249
2.6647 | -95.2112
2.6684 | -95.2793
2.6722 | -95.2677
2.676 | -95.256
2.6798 | -95.2936
2.6835 | -95.3835
2.6873 | -95.4599
2.6911 | -95.4478
2.6949 | -95.3345
2.6987 | -95.1596
2.7024 | -94.9974
2.7062 | -94.8987
2.71 | -94.8847
2.7138 | -94.9717
2.7176 | -95.1104
2.7213 | -95.2394
2.7251 | -95.1841
2.7289 | -95.131
2.7327 | -95.038
2.7365 | -94.9293
2.7402 | -94.8505
2.744 | -94.8424
2.7478 | -94.9072
2.7516 | -95.0347
2.7554 | -95.188
2.7591 | -95.3312
2.7629 | -95.4369
2.7667 | -95.47
2.7705 | -95.4674
2.7743 | -95.4422
2.778 | -95.4033
2.7818 | -95.3683
2.7856 | -95.3438
2.7894 | -95.3203
2.7932 | -95.3203
2.7969 | -95.3455
2.8007 | -95.3874
2.8045 | -95.4328
2.8083 | -95.4575
2.8121 | -95.4778
2.8158 | -95.5181
2.8196 | -95.5644
2.8234 | -95.6374
2.8272 | -95.7201
2.831 | -95.8009
2.8347 | -95.8796
2.8385 | -95.9523
2.8423 | -95.9958
2.8461 | -96.0169
2.8499 | -96.0008
2.8536 | -95.9291
2.8574 | -95.8201
2.8612 | -95.6686
2.865 | -95.4983
2.8688 | -95.3423
2.8725 | -95.2182
2.8763 | -95.1508
2.8801 | -95.1314
2.8839 | -95.1604
2.8876 | -95.2253
2.8914 | -95.3104
2.8952 | -95.3956
2.899 | -95.4677
2.9028 | -95.5323
2.9065 | -95.5946
2.9103 | -95.6729
2.9141 | -95.7609
2.9179 | -95.8404
2.9217 | -95.8875
2.9254 | -95.8952
2.9292 | -95.8685
2.933 | -95.825
2.9368 | -95.7891
2.9406 | -95.7949
2.9443 | -95.8516
2.9481 | -95.9629
2.9519 | -96.1047
2.9557 | -96.2561
2.9595 | -96.4061
2.9632 | -96.5403
2.967 | -96.6481
2.9708 | -96.7196
2.9746 | -96.7411
2.9784 | -96.7263
2.9821 | -96.6779
2.9859 | -96.6102
2.9897 | -96.5295
2.9935 | -96.4341
2.9973 | -96.3281
3.001 | -96.2141
3.0048 | -96.1108
3.0086 | -96.0385
3.0124 | -96.0074
3.0162 | -95.9922
3.0199 | -95.9716
3.0237 | -95.9459
3.0275 | -95.9077
3.0313 | -95.8638
3.0351 | -95.8183
3.0388 | -95.781
3.0426 | -95.7594
3.0464 | -95.756
3.0502 | -95.77
3.054 | -95.7847
3.0577 | -95.7876
3.0615 | -95.7804
3.0653 | -95.7738
3.0691 | -95.7677
3.0729 | -95.7645
3.0766 | -95.7866
3.0804 | -95.8529
3.0842 | -95.9607
3.088 | -96.0948
3.0918 | -96.238
3.0955 | -96.3526
3.0993 | -96.4126
3.1031 | -96.4177
3.1069 | -96.3638
3.1106 | -96.2595
3.1144 | -96.1418
3.1182 | -96.0495
3.122 | -96.0199
3.1258 | -96.0747
3.1295 | -96.188
3.1333 | -96.3319
3.1371 | -96.4866
3.1409 | -96.6331
3.1447 | -96.7567
3.1484 | -96.8445
3.1522 | -96.8667
3.156 | -96.8311
3.1598 | -96.7388
3.1636 | -96.6098
3.1673 | -96.4835
3.1711 | -96.36
3.1749 | -96.2383
3.1787 | -96.1436
3.1825 | -96.0822
3.1862 | -96.0668
3.19 | -96.1038
3.1938 | -96.206
3.1976 | -96.3502
3.2014 | -96.5014
3.2051 | -96.6178
3.2089 | -96.6775
3.2127 | -96.6954
3.2165 | -96.6938
3.2203 | -96.6959
3.224 | -96.6831
3.2278 | -96.6402
3.2316 | -96.5776
3.2354 | -96.5066
3.2392 | -96.4645
3.2429 | -96.4397
3.2467 | -96.4092
3.2505 | -96.3653
3.2543 | -96.2853
3.2581 | -96.1645
3.2618 | -96.0273
3.2656 | -95.9166
3.2694 | -95.9909
3.2732 | -96.0606
3.277 | -96.1884
3.2807 | -96.357
3.2845 | -96.5421
3.2883 | -96.7044
3.2921 | -96.8219
3.2959 | -96.9019
3.2996 | -96.9622
3.3034 | -96.9834
3.3072 | -96.9739
3.311 | -96.9463
3.3148 | -96.8913
3.3185 | -96.8143
3.3223 | -96.7249
3.3261 | -96.6523
3.3299 | -96.6093
3.3336 | -96.6151
3.3374 | -96.6705
3.3412 | -96.7653
3.345 | -96.897
3.3488 | -97.0455
3.3525 | -97.1927
3.3563 | -97.3131
3.3601 | -97.4116
3.3639 | -97.4589
3.3677 | -97.4493
3.3714 | -97.4012
3.3752 | -97.3437
3.379 | -97.2876
3.3828 | -97.2433
3.3866 | -97.2354
3.3903 | -97.2461
3.3941 | -97.2308
3.3979 | -97.1904
3.4017 | -97.12
3.4055 | -97.0103
3.4092 | -96.8886
3.413 | -96.7819
3.4168 | -96.6872
3.4206 | -96.6254
3.4244 | -96.5982
3.4281 | -96.6155
3.4319 | -96.6689
3.4357 | -96.7418
3.4395 | -96.8256
3.4433 | -96.8983
3.447 | -96.9511
3.4508 | -96.99
3.4546 | -97.0048
3.4584 | -97.0085
3.4622 | -97.0077
3.4659 | -97.0171
3.4697 | -97.0461
3.4735 | -97.113
3.4773 | -97.2027
3.4811 | -97.2929
3.4848 | -97.3532
3.4886 | -97.3757
3.4924 | -97.3605
3.4962 | -97.2837
3.5 | -97.1626
3.5037 | -97.0198
3.5075 | -96.8685
3.5113 | -96.7307
3.5151 | -96.6219
3.5189 | -96.5678
3.5226 | -96.5616
3.5264 | -96.6056
3.5302 | -96.6842
3.534 | -96.7859
3.5377 | -96.8982
3.5415 | -97.0151
3.5453 | -97.1265
3.5491 | -97.23
3.5529 | -97.3329
3.5566 | -97.4271
3.5604 | -97.5193
3.5642 | -97.6205
3.568 | -97.7324
3.5718 | -97.8398
3.5755 | -97.9308
3.5793 | -97.9748
3.5831 | -97.9569
3.5869 | -97.8946
3.5907 | -97.8008
3.5944 | -97.7032
3.5982 | -97.6196
3.602 | -97.5646
3.6058 | -97.5378
3.6096 | -97.5319
3.6133 | -97.5407
3.6171 | -97.5594
3.6209 | -97.5787
3.6247 | -97.5965
3.6285 | -97.6068
3.6322 | -97.6043
3.636 | -97.6093
3.6398 | -97.619
3.6436 | -97.6348
3.6474 | -97.6652
3.6511 | -97.7071
3.6549 | -97.756
3.6587 | -97.8054
3.6625 | -97.8612
3.6663 | -97.9141
3.67 | -97.9361
3.6738 | -97.9209
3.6776 | -97.8681
3.6814 | -97.7959
3.6852 | -97.7218
3.6889 | -97.6581
3.6927 | -97.6156
3.6965 | -97.605
3.7003 | -97.6139
3.7041 | -97.621
3.7078 | -97.6186
3.7116 | -97.6057
3.7154 | -97.5754
3.7192 | -97.5441
3.723 | -97.5373
3.7267 | -97.552
3.7305 | -97.5683
3.7343 | -97.5908
3.7381 | -97.6116
3.7418 | -97.6332
3.7456 | -97.6564
3.7494 | -97.68
3.7532 | -97.6859
3.757 | -97.6682
3.7607 | -97.6312
3.7645 | -97.5799
3.7683 | -97.5364
3.7721 | -97.5077
3.7759 | -97.4942
3.7796 | -97.4976
3.7834 | -97.5337
3.7872 | -97.6006
3.791 | -97.6901
3.7948 | -97.7941
3.7985 | -97.8919
3.8023 | -97.9648
3.8061 | -98.0131
3.8099 | -98.0487
3.8137 | -98.0795
3.8174 | -98.1107
3.8212 | -98.128
3.825 | -98.1346
3.8288 | -98.1255
3.8326 | -98.1031
3.8363 | -98.0656
3.8401 | -98.0168
3.8439 | -97.9598
3.8477 | -97.8957
3.8515 | -97.828
3.8552 | -97.7623
3.859 | -97.7084
3.8628 | -97.6699
3.8666 | -97.6516
3.8704 | -97.6653
3.8741 | -97.7094
3.8779 | -97.7794
3.8817 | -97.8741
3.8855 | -97.973
3.8893 | -98.0499
3.893 | -98.0976
3.8968 | -98.1227
3.9006 | -98.1371
3.9044 | -98.1562
3.9082 | -98.1875
3.9119 | -98.2478
3.9157 | -98.3233
3.9195 | -98.4067
3.9233 | -98.4821
3.9271 | -98.5417
3.9308 | -98.5693
3.9346 | -98.5582
3.9384 | -98.5135
3.9422 | -98.4373
3.946 | -98.327
3.9497 | -98.1961
3.9535 | -98.0619
3.9573 | -97.9389
3.9611 | -97.8443
3.9648 | -97.7858
3.9686 | -97.7691
3.9724 | -97.7978
3.9762 | -97.8493
3.98 | -97.9215
3.9837 | -97.998
3.9875 | -98.0702
3.9913 | -98.1253
3.9951 | -98.157
3.9989 | -98.1737
4.0026 | -98.1847
4.0064 | -98.1934
4.0102 | -98.216
4.014 | -98.2619
4.0178 | -98.3366
4.0215 | -98.4319
4.0253 | -98.531
4.0291 | -98.6231
4.0329 | -98.7041
4.0367 | -98.7421
4.0404 | -98.7441
4.0442 | -98.7097
4.048 | -98.6555
4.0518 | -98.5872
4.0556 | -98.515
4.0593 | -98.4624
4.0631 | -98.4395
4.0669 | -98.442
4.0707 | -98.4729
4.0745 | -98.5162
4.0782 | -98.5513
4.082 | -98.5634
4.0858 | -98.5459
4.0896 | -98.5015
4.0934 | -98.436
4.0971 | -98.3572
4.1009 | -98.2911
4.1047 | -98.2439
4.1085 | -98.2235
4.1123 | -98.2231
4.116 | -98.2483
4.1198 | -98.2883
4.1236 | -98.3362
4.1274 | -98.3817
4.1312 | -98.4215
4.1349 | -98.4482
4.1387 | -98.4582
4.1425 | -98.4503
4.1463 | -98.4161
4.1501 | -98.3675
4.1538 | -98.316
4.1576 | -98.2691
4.1614 | -98.2338
4.1652 | -98.2253
4.1689 | -98.2338
4.1727 | -98.2575
4.1765 | -98.3087
4.1803 | -98.3825
4.1841 | -98.4729
4.1878 | -98.5665
4.1916 | -98.6505
4.1954 | -98.7327
4.1992 | -98.798
4.203 | -98.8203
4.2067 | -98.801
4.2105 | -98.751
4.2143 | -98.6734
4.2181 | -98.5845
4.2219 | -98.5085
4.2256 | -98.4615
4.2294 | -98.447
4.2332 | -98.4741
4.237 | -98.538
4.2408 | -98.6205
4.2445 | -98.715
4.2483 | -98.8099
4.2521 | -98.8933
4.2559 | -98.9691
4.2597 | -99.0225
4.2634 | -99.0647
4.2672 | -99.1156
4.271 | -99.1713
4.2748 | -99.235
4.2786 | -99.3065
4.2823 | -99.3814
4.2861 | -99.4562
4.2899 | -99.5176
4.2937 | -99.5515
4.2975 | -99.5538
4.3012 | -99.5153
4.305 | -99.4532
4.3088 | -99.3722
4.3126 | -99.2878
4.3164 | -99.2114
4.3201 | -99.1321
4.3239 | -99.0723
4.3277 | -99.0326
4.3315 | -99.0091
4.3353 | -99.0139
4.339 | -99.0429
4.3428 | -99.0938
4.3466 | -99.1615
4.3504 | -99.2271
4.3542 | -99.2797
4.3579 | -99.3136
4.3617 | -99.2878
4.3655 | -99.2812
4.3693 | -99.2622
4.3731 | -99.2302
4.3768 | -99.194
4.3806 | -99.1614
4.3844 | -99.1366
4.3882 | -99.1287
4.3919 | -99.1148
4.3957 | -99.0963
4.3995 | -99.0796
4.4033 | -99.064
4.4071 | -99.0569
4.4108 | -99.0658
4.4146 | -99.083
4.4184 | -99.1136
4.4222 | -99.1519
4.426 | -99.1859
4.4297 | -99.2063
4.4335 | -99.2217
4.4373 | -99.227
4.4411 | -99.2301
4.4449 | -99.2278
4.4486 | -99.2286
4.4524 | -99.2282
4.4562 | -99.2158
4.46 | -99.2019
4.4638 | -99.197
4.4675 | -99.1904
4.4713 | -99.1816
4.4751 | -99.1698
4.4789 | -99.1588
4.4827 | -99.1496
4.4864 | -99.144
4.4902 | -99.1471
4.494 | -99.1525
4.4978 | -99.1584
4.5016 | -99.1677
4.5053 | -99.172
4.5091 | -99.1722
4.5129 | -99.1789
4.5167 | -99.1882
4.5205 | -99.1979
4.5242 | -99.2083
4.528 | -99.2206
4.5318 | -99.2227
4.5356 | -99.2139
4.5394 | -99.2054
4.5431 | -99.193
4.5469 | -99.1879
4.5507 | -99.1984
4.5545 | -99.2255
4.5583 | -99.2737
4.562 | -99.3399
4.5658 | -99.4151
4.5696 | -99.4832
4.5734 | -99.539
4.5772 | -99.5792
4.5809 | -99.6028
4.5847 | -99.6106
4.5885 | -99.6203
4.5923 | -99.6311
4.596 | -99.6513
4.5998 | -99.6817
4.6036 | -99.7055
4.6074 | -99.7166
4.6112 | -99.7182
4.6149 | -99.7036
4.6187 | -99.672
4.6225 | -99.628
4.6263 | -99.5769
4.6301 | -99.5254
4.6338 | -99.4772
4.6376 | -99.452
4.6414 | -99.4607
4.6452 | -99.4968
4.649 | -99.564
4.6527 | -99.6568
4.6565 | -99.7646
4.6603 | -99.8705
4.6641 | -99.9671
4.6679 | -100.038
4.6716 | -100.0813
4.6754 | -100.1042
4.6792 | -100.0936
4.683 | -100.0485
4.6868 | -99.9815
4.6905 | -99.8976
4.6943 | -99.8143
4.6981 | -99.7378
4.7019 | -99.6743
4.7057 | -99.6302
4.7094 | -99.6081
4.7132 | -99.6052
4.717 | -99.6198
4.7208 | -99.6439
4.7246 | -99.6679
4.7283 | -99.691
4.7321 | -99.7198
4.7359 | -99.738
4.7397 | -99.7442
4.7435 | -99.7376
4.7472 | -99.7213
4.751 | -99.704
4.7548 | -99.6824
4.7586 | -99.661
4.7624 | -99.647
4.7661 | -99.6431
4.7699 | -99.661
4.7737 | -99.701
4.7775 | -99.7585
4.7813 | -99.8311
4.785 | -99.9064
4.7888 | -99.9731
4.7926 | -100.0175
4.7964 | -100.0422
4.8002 | -100.0414
4.8039 | -100.0179
4.8077 | -99.9831
4.8115 | -99.9474
4.8153 | -99.9154
4.819 | -99.8803
4.8228 | -99.8443
4.8266 | -99.8131
4.8304 | -99.7933
4.8342 | -99.7761
4.8379 | -99.7747
4.8417 | -99.7936
4.8455 | -99.8255
4.8493 | -99.8674
4.8531 | -99.9089
4.8568 | -99.9475
4.8606 | -99.974
4.8644 | -99.9881
4.8682 | -99.9912
4.872 | -99.9923
4.8757 | -100.0041
4.8795 | -100.0231
4.8833 | -100.0536
4.8871 | -100.0822
4.8909 | -100.12
4.8946 | -100.1576
4.8984 | -100.1863
4.9022 | -100.2023
4.906 | -100.1918
4.9098 | -100.1953
4.9135 | -100.1928
4.9173 | -100.1824
4.9211 | -100.1771
4.9249 | -100.1765
4.9287 | -100.1802
4.9324 | -100.1853
4.9362 | -100.193
4.94 | -100.1964
4.9438 | -100.194
4.9476 | -100.1802
4.9513 | -100.1557
4.9551 | -100.1256
4.9589 | -100.0914
4.9627 | -100.0585
4.9665 | -100.0257
4.9702 | -99.997
4.974 | -99.9804
4.9778 | -99.9792
4.9816 | -99.9823
4.9854 | -99.9912
4.9891 | -100.0101
4.9929 | -100.0322
4.9967 | -100.054
5.0005 | -100.0801
5.0043 | -100.1156
5.008 | -100.1578
5.0118 | -100.2035
5.0156 | -100.2546
5.0194 | -100.3119
5.0232 | -100.3598
5.0269 | -100.3979
5.0307 | -100.4339
5.0345 | -100.4595
5.0383 | -100.4754
5.042 | -100.479
5.0458 | -100.4654
5.0496 | -100.4373
5.0534 | -100.3926
5.0572 | -100.3271
5.0609 | -100.2483
5.0647 | -100.1547
5.0685 | -100.0443
5.0723 | -99.9299
5.0761 | -99.8193
5.0798 | -99.7076
5.0836 | -99.6161
5.0874 | -99.5574
5.0912 | -99.5326
5.095 | -99.5409
5.0987 | -99.5847
5.1025 | -99.6654
5.1063 | -99.7761
5.1101 | -99.9007
5.1139 | -100.0373
5.1176 | -100.1745
5.1214 | -100.3051
5.1252 | -100.4227
5.129 | -100.5271
5.1328 | -100.6041
5.1365 | -100.6654
5.1403 | -100.7155
5.1441 | -100.7568
5.1479 | -100.7942
5.1517 | -100.8279
5.1554 | -100.8646
5.1592 | -100.9021
5.163 | -100.9219
5.1668 | -100.921
5.1706 | -100.9072
5.1743 | -100.8823
5.1781 | -100.8494
5.1819 | -100.8142
5.1857 | -100.7761
5.1895 | -100.742
5.1932 | -100.7113
5.197 | -100.6888
5.2008 | -100.6882
5.2046 | -100.6976
5.2084 | -100.718
5.2121 | -100.7507
5.2159 | -100.7789
5.2197 | -100.8
5.2235 | -100.8063
5.2273 | -100.8036
5.231 | -100.7941
5.2348 | -100.7752
5.2386 | -100.7547
5.2424 | -100.7333
5.2461 | -100.7221
5.2499 | -100.7095
5.2537 | -100.6886
5.2575 | -100.6615
5.2613 | -100.6363
5.265 | -100.6202
5.2688 | -100.6038
5.2726 | -100.5992
5.2764 | -100.5964
5.2802 | -100.5967
5.2839 | -100.602
5.2877 | -100.6022
5.2915 | -100.6024
5.2953 | -100.5988
5.2991 | -100.5883
5.3028 | -100.5772
5.3066 | -100.5641
5.3104 | -100.5516
5.3142 | -100.5374
5.318 | -100.5224
5.3217 | -100.5048
5.3255 | -100.4797
5.3293 | -100.4544
5.3331 | -100.4262
5.3369 | -100.4052
5.3406 | -100.3899
5.3444 | -100.3951
5.3482 | -100.425
5.352 | -100.4748
5.3558 | -100.5475
5.3595 | -100.628
5.3633 | -100.7166
5.3671 | -100.8076
5.3709 | -100.8899
5.3747 | -100.9551
5.3784 | -100.9957
5.3822 | -101.0137
5.386 | -101.0037
5.3898 | -100.9735
5.3936 | -100.932
5.3973 | -100.8738
5.4011 | -100.8142
5.4049 | -100.76
5.4087 | -100.7115
5.4125 | -100.6801
5.4162 | -100.6589
5.42 | -100.6434
5.4238 | -100.629
5.4276 | -100.6267
5.4314 | -100.6293
5.4351 | -100.6305
5.4389 | -100.6383
5.4427 | -100.6501
5.4465 | -100.6662
5.4502 | -100.6883
5.454 | -100.7038
5.4578 | -100.7259
5.4616 | -100.7515
5.4654 | -100.7694
5.4691 | -100.7848
5.4729 | -100.8044
5.4767 | -100.8249
5.4805 | -100.8426
5.4843 | -100.8605
5.488 | -100.8722
5.4918 | -100.8862
5.4956 | -100.8978
5.4994 | -100.9073
5.5032 | -100.913
5.5069 | -100.9119
5.5107 | -100.8974
5.5145 | -100.8811
5.5183 | -100.8696
5.5221 | -100.8542
5.5258 | -100.8362
5.5296 | -100.825
5.5334 | -100.8211
5.5372 | -100.8231
5.541 | -100.8248
5.5447 | -100.8262
5.5485 | -100.8268
5.5523 | -100.8355
5.5561 | -100.8481
5.5599 | -100.8489
5.5636 | -100.8413
5.5674 | -100.8305
5.5712 | -100.8075
5.575 | -100.7815
5.5788 | -100.7441
5.5825 | -100.6954
5.5863 | -100.6482
5.5901 | -100.6086
5.5939 | -100.5749
5.5977 | -100.552
5.6014 | -100.5364
5.6052 | -100.5356
5.609 | -100.5537
5.6128 | -100.5847
5.6166 | -100.6228
5.6203 | -100.6679
5.6241 | -100.7157
5.6279 | -100.7656
5.6317 | -100.8155
5.6355 | -100.8548
5.6392 | -100.8796
5.643 | -100.9023
5.6468 | -100.9254
5.6506 | -100.9385
5.6544 | -100.9389
5.6581 | -100.933
5.6619 | -100.9149
5.6657 | -100.8934
5.6695 | -100.874
5.6733 | -100.8531
5.677 | -100.8421
5.6808 | -100.8435
5.6846 | -100.8597
5.6884 | -100.8884
5.6921 | -100.9239
5.6959 | -100.9664
5.6997 | -101.0151
5.7035 | -101.0667
5.7073 | -101.1097
5.711 | -101.1345
5.7148 | -101.1395
5.7186 | -101.1274
5.7224 | -101.1071
5.7262 | -101.0761
5.7299 | -101.0283
5.7337 | -100.985
5.7375 | -100.9528
5.7413 | -100.9324
5.7451 | -100.9261
5.7488 | -100.9384
5.7526 | -100.9646
5.7564 | -101.0101
5.7602 | -101.0636
5.764 | -101.1189
5.7677 | -101.1658
5.7715 | -101.206
5.7753 | -101.2385
5.7791 | -101.2532
5.7829 | -101.2502
5.7866 | -101.2344
5.7904 | -101.2006
5.7942 | -101.1568
5.798 | -101.1065
5.8018 | -101.0469
5.8055 | -100.9738
5.8093 | -100.9008
5.8131 | -100.8349
5.8169 | -100.7763
5.8207 | -100.7206
5.8244 | -100.6733
5.8282 | -100.6374
5.832 | -100.6086
5.8358 | -100.5828
5.8396 | -100.5665
5.8433 | -100.5485
5.8471 | -100.5377
5.8509 | -100.5306
5.8547 | -100.531
5.8585 | -100.5267
5.8622 | -100.5229
5.866 | -100.5254
5.8698 | -100.5286
5.8736 | -100.5364
5.8773 | -100.549
5.8811 | -100.5654
5.8849 | -100.5809
5.8887 | -100.5995
5.8925 | -100.6244
5.8962 | -100.6475
5.9 | -100.6615
5.9038 | -100.6698
5.9076 | -100.6696
5.9114 | -100.6662
5.9151 | -100.6639
5.9189 | -100.6706
5.9227 | -100.6838
5.9265 | -100.6978
5.9303 | -100.7158
5.934 | -100.743
5.9378 | -100.776
5.9416 | -100.8106
5.9454 | -100.8493
5.9492 | -100.8887
5.9529 | -100.9276
5.9567 | -100.9737
5.9605 | -101.0244
5.9643 | -101.0825
5.9681 | -101.143
5.9718 | -101.2135
5.9756 | -101.2772
5.9794 | -101.3375
5.9832 | -101.3965
5.987 | -101.4555
5.9907 | -101.5124
5.9945 | -101.5123
5.9983 | -101.5308
6.0021 | -101.5291
6.0059 | -101.5061
6.0096 | -101.4678
6.0134 | -101.4171
6.0172 | -101.3571
6.021 | -101.2964
6.0248 | -101.228
6.0285 | -101.1671
6.0323 | -101.1195
6.0361 | -101.0882
6.0399 | -101.075
6.0437 | -101.0807
6.0474 | -101.1011
6.0512 | -101.1347
6.055 | -101.1867
6.0588 | -101.2494
6.0626 | -101.3169
6.0663 | -101.388
6.0701 | -101.4526
6.0739 | -101.5137
6.0777 | -101.5655
6.0815 | -101.6045
6.0852 | -101.6329
6.089 | -101.6483
6.0928 | -101.6404
6.0966 | -101.6251
6.1003 | -101.6035
6.1041 | -101.5792
6.1079 | -101.5608
6.1117 | -101.5358
6.1155 | -101.5064
6.1192 | -101.4876
6.123 | -101.4688
6.1268 | -101.4539
6.1306 | -101.4429
6.1344 | -101.4293
6.1381 | -101.4191
6.1419 | -101.4081
6.1457 | -101.3966
6.1495 | -101.3927
6.1533 | -101.3971
6.157 | -101.3969
6.1608 | -101.4029
6.1646 | -101.4166
6.1684 | -101.4222
6.1722 | -101.4265
6.1759 | -101.4289
6.1797 | -101.4269
6.1835 | -101.4204
6.1873 | -101.4162
6.1911 | -101.3986
6.1948 | -101.3782
6.1986 | -101.3589
6.2024 | -101.3525
6.2062 | -101.3654
6.21 | -101.3885
6.2137 | -101.4146
6.2175 | -101.4535
6.2213 | -101.5005
6.2251 | -101.548
6.2289 | -101.5993
6.2326 | -101.6481
6.2364 | -101.686
6.2402 | -101.7154
6.244 | -101.7288
6.2478 | -101.7334
6.2515 | -101.7364
6.2553 | -101.7332
6.2591 | -101.7199
6.2629 | -101.7056
6.2667 | -101.6871
6.2704 | -101.6712
6.2742 | -101.6531
6.278 | -101.6367
6.2818 | -101.6193
6.2856 | -101.6127
6.2893 | -101.6089
6.2931 | -101.6055
6.2969 | -101.5946
6.3007 | -101.5871
6.3045 | -101.5838
6.3082 | -101.5841
6.312 | -101.5836
6.3158 | -101.5871
6.3196 | -101.5931
6.3233 | -101.5955
6.3271 | -101.6001
6.3309 | -101.6107
6.3347 | -101.6152
6.3385 | -101.6122
6.3422 | -101.6063
6.346 | -101.5964
6.3498 | -101.5803
6.3536 | -101.5666
6.3574 | -101.5566
6.3611 | -101.5498
6.3649 | -101.5465
6.3687 | -101.5591
6.3725 | -101.585
6.3763 | -101.6204
6.38 | -101.6593
6.3838 | -101.7055
6.3876 | -101.7547
6.3914 | -101.8011
6.3952 | -101.8364
6.3989 | -101.8651
6.4027 | -101.8885
6.4065 | -101.9095
6.4103 | -101.9219
6.4141 | -101.9359
6.4178 | -101.9515
6.4216 | -101.9626
6.4254 | -101.9781
6.4292 | -101.9971
6.433 | -102.0195
6.4367 | -102.0369
6.4405 | -102.0586
6.4443 | -102.0765
6.4481 | -102.084
6.4519 | -102.0837
6.4556 | -102.0799
6.4594 | -102.0637
6.4632 | -102.037
6.467 | -102.0017
6.4708 | -101.9702
6.4745 | -101.9446
6.4783 | -101.9273
6.4821 | -101.9233
6.4859 | -101.9351
6.4897 | -101.9568
6.4934 | -101.991
6.4972 | -102.0392
6.501 | -102.0924
6.5048 | -102.1482
6.5086 | -102.2
6.5123 | -102.2467
6.5161 | -102.2874
6.5199 | -102.3123
6.5237 | -102.3249
6.5274 | -102.3204
6.5312 | -102.3008
6.535 | -102.2651
6.5388 | -102.2182
6.5426 | -102.1814
6.5463 | -102.1475
6.5501 | -102.1155
6.5539 | -102.0889
6.5577 | -102.0753
6.5615 | -102.0698
6.5652 | -102.0689
6.569 | -102.0726
6.5728 | -102.0848
6.5766 | -102.0973
6.5804 | -102.1078
6.5841 | -102.1267
6.5879 | -102.156
6.5917 | -102.1873
6.5955 | -102.2174
6.5993 | -102.2473
6.603 | -102.2724
6.6068 | -102.2913
6.6106 | -102.3022
6.6144 | -102.3057
6.6182 | -102.3039
6.6219 | -102.2951
6.6257 | -102.2884
6.6295 | -102.2782
6.6333 | -102.2652
6.6371 | -102.2569
6.6408 | -102.2416
6.6446 | -102.2232
6.6484 | -102.2104
6.6522 | -102.1979
6.656 | -102.1811
6.6597 | -102.1624
6.6635 | -102.1474
6.6673 | -102.1332
6.6711 | -102.1155
6.6749 | -102.0979
6.6786 | -102.0846
6.6824 | -102.0795
6.6862 | -102.0737
6.69 | -102.0643
6.6938 | -102.0611
6.6975 | -102.0639
6.7013 | -102.0779
6.7051 | -102.1011
6.7089 | -102.134
6.7127 | -102.1739
6.7164 | -102.2137
6.7202 | -102.2525
6.724 | -102.2886
6.7278 | -102.3127
6.7316 | -102.3303
6.7353 | -102.3328
6.7391 | -102.3156
6.7429 | -102.2901
6.7467 | -102.2575
6.7504 | -102.2129
6.7542 | -102.1632
6.758 | -102.1123
6.7618 | -102.0657
6.7656 | -102.0185
6.7693 | -101.9763
6.7731 | -101.9401
6.7769 | -101.9097
6.7807 | -101.8831
6.7845 | -101.8598
6.7882 | -101.843
6.792 | -101.8442
6.7958 | -101.8559
6.7996 | -101.8768
6.8034 | -101.9077
6.8071 | -101.9515
6.8109 | -102.0013
6.8147 | -102.0574
6.8185 | -102.1279
6.8223 | -102.203
6.826 | -102.2829
6.8298 | -102.366
6.8336 | -102.4454
6.8374 | -102.5184
6.8412 | -102.5834
6.8449 | -102.6287
6.8487 | -102.6618
6.8525 | -102.6827
6.8563 | -102.6838
6.8601 | -102.6735
6.8638 | -102.648
6.8676 | -102.6152
6.8714 | -102.5795
6.8752 | -102.5453
6.879 | -102.5156
6.8827 | -102.4944
6.8865 | -102.4771
6.8903 | -102.4677
6.8941 | -102.4645
6.8979 | -102.4681
6.9016 | -102.4778
6.9054 | -102.4895
6.9092 | -102.5076
6.913 | -102.5323
6.9168 | -102.5604
6.9205 | -102.5778
6.9243 | -102.5968
6.9281 | -102.6123
6.9319 | -102.6194
6.9357 | -102.6193
6.9394 | -102.6127
6.9432 | -102.6017
6.947 | -102.5881
6.9508 | -102.5695
6.9546 | -102.555
6.9583 | -102.5442
6.9621 | -102.5313
6.9659 | -102.523
6.9697 | -102.5244
6.9734 | -102.5294
6.9772 | -102.5292
6.981 | -102.5288
6.9848 | -102.5331
6.9886 | -102.5444
6.9923 | -102.5554
6.9961 | -102.5642
6.9999 | -102.5689
7.0037 | -102.5756
7.0075 | -102.5827
7.0112 | -102.5844
7.015 | -102.5806
7.0188 | -102.5778
7.0226 | -102.581
7.0264 | -102.5802
7.0301 | -102.5688
7.0339 | -102.5571
7.0377 | -102.5491
7.0415 | -102.5453
7.0453 | -102.5422
7.049 | -102.5408
7.0528 | -102.5425
7.0566 | -102.5434
7.0604 | -102.5487
7.0642 | -102.5542
7.0679 | -102.5563
7.0717 | -102.5553
7.0755 | -102.5507
7.0793 | -102.5479
7.0831 | -102.5451
7.0868 | -102.5261
7.0906 | -102.5152
7.0944 | -102.502
7.0982 | -102.4916
7.102 | -102.4844
7.1057 | -102.4835
7.1095 | -102.4852
7.1133 | -102.4941
7.1171 | -102.5138
7.1209 | -102.5474
7.1246 | -102.5854
7.1284 | -102.6297
7.1322 | -102.6828
7.136 | -102.7355
7.1398 | -102.785
7.1435 | -102.8289
7.1473 | -102.8675
7.1511 | -102.8918
7.1549 | -102.903
7.1587 | -102.903
7.1624 | -102.8931
7.1662 | -102.8685
7.17 | -102.8274
7.1738 | -102.7842
7.1775 | -102.7355
7.1813 | -102.6863
7.1851 | -102.6337
7.1889 | -102.5778
7.1927 | -102.529
7.1964 | -102.4833
7.2002 | -102.4347
7.204 | -102.3838
7.2078 | -102.3382
7.2116 | -102.292
7.2153 | -102.2473
7.2191 | -102.2058
7.2229 | -102.1622
7.2267 | -102.1239
7.2305 | -102.0947
7.2342 | -102.0784
7.238 | -102.0728
7.2418 | -102.0788
7.2456 | -102.1043
7.2494 | -102.1415
7.2531 | -102.1885
7.2569 | -102.247
7.2607 | -102.3182
7.2645 | -102.393
7.2683 | -102.4721
7.272 | -102.5551
7.2758 | -102.642
7.2796 | -102.7277
7.2834 | -102.7906
7.2872 | -102.8427
7.2909 | -102.8819
7.2947 | -102.9099
7.2985 | -102.9229
7.3023 | -102.9169
7.3061 | -102.8973
7.3098 | -102.8695
7.3136 | -102.8325
7.3174 | -102.7823
7.3212 | -102.7349
7.325 | -102.694
7.3287 | -102.6608
7.3325 | -102.6387
7.3363 | -102.6228
7.3401 | -102.6165
7.3439 | -102.6235
7.3476 | -102.6434
7.3514 | -102.6751
7.3552 | -102.7106
7.359 | -102.7504
7.3628 | -102.7902
7.3665 | -102.8249
7.3703 | -102.8594
7.3741 | -102.892
7.3779 | -102.9118
7.3816 | -102.9194
7.3854 | -102.9149
7.3892 | -102.9059
7.393 | -102.8992
7.3968 | -102.8827
7.4005 | -102.8658
7.4043 | -102.8478
7.4081 | -102.8352
7.4119 | -102.8262
7.4157 | -102.8283
7.4194 | -102.8411
7.4232 | -102.8639
7.427 | -102.8972
7.4308 | -102.9374
7.4346 | -102.9767
7.4383 | -103.0189
7.4421 | -103.0624
7.4459 | -103.1042
7.4497 | -103.1464
7.4535 | -103.1829
7.4572 | -103.2055
7.461 | -103.217
7.4648 | -103.2242
7.4686 | -103.2258
7.4724 | -103.2151
7.4761 | -103.1953
7.4799 | -103.1703
7.4837 | -103.1396
7.4875 | -103.1061
7.4913 | -103.0716
7.495 | -103.0393
7.4988 | -103.0109
7.5026 | -102.985
7.5064 | -102.9641
7.5102 | -102.9452
7.5139 | -102.9364
7.5177 | -102.9356
7.5215 | -102.9477
7.5253 | -102.9602
7.5291 | -102.974
7.5328 | -102.9937
7.5366 | -103.0152
7.5404 | -103.0333
7.5442 | -103.0561
7.548 | -103.076
7.5517 | -103.0907
7.5555 | -103.098
7.5593 | -103.0925
7.5631 | -103.0806
7.5669 | -103.0657
7.5706 | -103.0453
7.5744 | -103.0263
7.5782 | -103.002
7.582 | -102.9839
7.5858 | -102.9705
7.5895 | -102.9599
7.5933 | -102.9596
7.5971 | -102.9626
7.6009 | -102.969
7.6046 | -102.9769
7.6084 | -102.9856
7.6122 | -102.997
7.616 | -103.0148
7.6198 | -103.0345
7.6235 | -103.0462
7.6273 | -103.0572
7.6311 | -103.0774
7.6349 | -103.0967
7.6387 | -103.1257
7.6424 | -103.1554
7.6462 | -103.1813
7.65 | -103.2054
7.6538 | -103.2371
7.6576 | -103.2666
7.6613 | -103.2922
7.6651 | -103.3138
7.6689 | -103.335
7.6727 | -103.3493
7.6765 | -103.3586
7.6802 | -103.36
7.684 | -103.3552
7.6878 | -103.3397
7.6916 | -103.3141
7.6954 | -103.2823
7.6991 | -103.2484
7.7029 | -103.2093
7.7067 | -103.1667
7.7105 | -103.1305
7.7143 | -103.0936
7.718 | -103.0582
7.7218 | -103.0272
7.7256 | -102.9994
7.7294 | -102.9761
7.7332 | -102.9617
7.7369 | -102.9579
7.7407 | -102.9618
7.7445 | -102.9732
7.7483 | -102.9941
7.7521 | -103.0224
7.7558 | -103.0539
7.7596 | -103.097
7.7634 | -103.1494
7.7672 | -103.2064
7.771 | -103.2605
7.7747 | -103.3149
7.7785 | -103.3677
7.7823 | -103.4152
7.7861 | -103.4489
7.7899 | -103.4759
7.7936 | -103.4924
7.7974 | -103.5054
7.8012 | -103.5086
7.805 | -103.5039
7.8087 | -103.4979
7.8125 | -103.4857
7.8163 | -103.4628
7.8201 | -103.4406
7.8239 | -103.4194
7.8276 | -103.3973
7.8314 | -103.3806
7.8352 | -103.361
7.839 | -103.3434
7.8428 | -103.331
7.8465 | -103.3211
7.8503 | -103.3127
7.8541 | -103.3039
7.8579 | -103.2959
7.8617 | -103.2839
7.8654 | -103.277
7.8692 | -103.273
7.873 | -103.2708
7.8768 | -103.2691
7.8806 | -103.2692
7.8843 | -103.2681
7.8881 | -103.2662
7.8919 | -103.2654
7.8957 | -103.265
7.8995 | -103.2645
7.9032 | -103.2645
7.907 | -103.2608
7.9108 | -103.254
7.9146 | -103.2423
7.9184 | -103.2265
7.9221 | -103.2123
7.9259 | -103.2028
7.9297 | -103.1957
7.9335 | -103.1963
7.9373 | -103.1978
7.941 | -103.2009
7.9448 | -103.2109
7.9486 | -103.2322
7.9524 | -103.264
7.9562 | -103.297
7.9599 | -103.3324
7.9637 | -103.3727
7.9675 | -103.4182
7.9713 | -103.4666
7.9751 | -103.5055
7.9788 | -103.5448
7.9826 | -103.5699
7.9864 | -103.5924
7.9902 | -103.6116
7.994 | -103.6226
7.9977 | -103.6263
8.0015 | -103.6271
8.0053 | -103.6201
8.0091 | -103.6122
8.0129 | -103.6017
8.0166 | -103.5922
8.0204 | -103.5794
8.0242 | -103.5652
8.028 | -103.5514
8.0317 | -103.5359
8.0355 | -103.5227
8.0393 | -103.5037
8.0431 | -103.4816
8.0469 | -103.4634
8.0506 | -103.4465
8.0544 | -103.4297
8.0582 | -103.4114
8.062 | -103.39
8.0658 | -103.3618
8.0695 | -103.3361
8.0733 | -103.3072
8.0771 | -103.274
8.0809 | -103.2398
8.0847 | -103.2121
8.0884 | -103.1879
8.0922 | -103.1721
8.096 | -103.1533
8.0998 | -103.1393
8.1036 | -103.1294
8.1073 | -103.1295
8.1111 | -103.1365
8.1149 | -103.1494
8.1187 | -103.1643
8.1225 | -103.187
8.1262 | -103.2137
8.13 | -103.2443
8.1338 | -103.2714
8.1376 | -103.2984
8.1414 | -103.3272
8.1451 | -103.3552
8.1489 | -103.3834
8.1527 | -103.4135
8.1565 | -103.4389
8.1603 | -103.4573
8.164 | -103.4768
8.1678 | -103.499
8.1716 | -103.5241
8.1754 | -103.5552
8.1792 | -103.5929
8.1829 | -103.6249
8.1867 | -103.6584
8.1905 | -103.692
8.1943 | -103.7283
8.1981 | -103.7622
8.2018 | -103.7896
8.2056 | -103.806
8.2094 | -103.8202
8.2132 | -103.8281
8.217 | -103.829
8.2207 | -103.8175
8.2245 | -103.7967
8.2283 | -103.7664
8.2321 | -103.7284
8.2358 | -103.6933
8.2396 | -103.661
8.2434 | -103.6294
8.2472 | -103.5926
8.251 | -103.5581
8.2547 | -103.5318
8.2585 | -103.5088
8.2623 | -103.487
8.2661 | -103.4667
8.2699 | -103.4537
8.2736 | -103.4442
8.2774 | -103.4433
8.2812 | -103.4437
8.285 | -103.439
8.2888 | -103.4329
8.2925 | -103.4295
8.2963 | -103.4312
8.3001 | -103.4327
8.3039 | -103.4318
8.3077 | -103.4343
8.3114 | -103.4342
8.3152 | -103.4314
8.319 | -103.4282
8.3228 | -103.4271
8.3266 | -103.4271
8.3303 | -103.4269
8.3341 | -103.4261
8.3379 | -103.4296
8.3417 | -103.4415
8.3455 | -103.4544
8.3492 | -103.4641
8.353 | -103.4773
8.3568 | -103.4954
8.3606 | -103.5109
8.3644 | -103.5304
8.3681 | -103.5473
8.3719 | -103.5639
8.3757 | -103.5815
8.3795 | -103.5964
8.3833 | -103.6095
8.387 | -103.6194
8.3908 | -103.6296
8.3946 | -103.6369
8.3984 | -103.6404
8.4022 | -103.6416
8.4059 | -103.6426
8.4097 | -103.6473
8.4135 | -103.6546
8.4173 | -103.6586
8.4211 | -103.6609
8.4248 | -103.6676
8.4286 | -103.6768
8.4324 | -103.6787
8.4362 | -103.6815
8.44 | -103.6861
8.4437 | -103.6922
8.4475 | -103.6977
8.4513 | -103.7044
8.4551 | -103.7113
8.4588 | -103.724
8.4626 | -103.7385
8.4664 | -103.7556
8.4702 | -103.7811
8.474 | -103.8103
8.4777 | -103.8393
8.4815 | -103.8746
8.4853 | -103.9086
8.4891 | -103.9441
8.4929 | -103.9839
8.4966 | -104.019
8.5004 | -104.0472
8.5042 | -104.0744
8.508 | -104.097
8.5118 | -104.1176
8.5155 | -104.1359
8.5193 | -104.1488
8.5231 | -104.1582
8.5269 | -104.1702
8.5307 | -104.1811
8.5344 | -104.183
8.5382 | -104.1773
8.542 | -104.1756
8.5458 | -104.1659
8.5496 | -104.1484
8.5533 | -104.1287
8.5571 | -104.1116
8.5609 | -104.0922
8.5647 | -104.0628
8.5685 | -104.0279
8.5722 | -103.9882
8.576 | -103.9421
8.5798 | -103.8922
8.5836 | -103.839
8.5874 | -103.7859
8.5911 | -103.7343
8.5949 | -103.6789
8.5987 | -103.6268
8.6025 | -103.5823
8.6063 | -103.5379
8.61 | -103.5016
8.6138 | -103.4695
8.6176 | -103.4465
8.6214 | -103.4356
8.6252 | -103.4281
8.6289 | -103.4206
8.6327 | -103.4205
8.6365 | -103.4275
8.6403 | -103.4399
8.6441 | -103.4625
8.6478 | -103.492
8.6516 | -103.5191
8.6554 | -103.5508
8.6592 | -103.5893
8.663 | -103.6298
8.6667 | -103.666
8.6705 | -103.7044
8.6743 | -103.7418
8.6781 | -103.7759
8.6818 | -103.8061
8.6856 | -103.8328
8.6894 | -103.8563
8.6932 | -103.8699
8.697 | -103.8736
8.7007 | -103.8672
8.7045 | -103.8604
8.7083 | -103.8489
8.7121 | -103.8272
8.7159 | -103.8043
8.7196 | -103.7877
8.7234 | -103.7723
8.7272 | -103.7559
8.731 | -103.7379
8.7348 | -103.7225
8.7385 | -103.7067
8.7423 | -103.692
8.7461 | -103.679
8.7499 | -103.6671
8.7537 | -103.6563
8.7574 | -103.651
8.7612 | -103.6477
8.765 | -103.6458
8.7688 | -103.6473
8.7726 | -103.6539
8.7763 | -103.6633
8.7801 | -103.6797
8.7839 | -103.6991
8.7877 | -103.7197
8.7915 | -103.7476
8.7952 | -103.7767
8.799 | -103.8052
8.8028 | -103.8377
8.8066 | -103.8686
8.8104 | -103.8985
8.8141 | -103.9233
8.8179 | -103.9498
8.8217 | -103.9779
8.8255 | -104.004
8.8293 | -104.0293
8.833 | -104.046
8.8368 | -104.0562
8.8406 | -104.0673
8.8444 | -104.0748
8.8482 | -104.0792
8.8519 | -104.0813
8.8557 | -104.0855
8.8595 | -104.091
8.8633 | -104.0962
8.8671 | -104.0978
8.8708 | -104.1026
8.8746 | -104.1065
8.8784 | -104.1086
8.8822 | -104.1124
8.8859 | -104.122
8.8897 | -104.1297
8.8935 | -104.1358
8.8973 | -104.1423
8.9011 | -104.1526
8.9048 | -104.1608
8.9086 | -104.1695
8.9124 | -104.1678
8.9162 | -104.1517
8.92 | -104.1327
8.9237 | -104.1083
8.9275 | -104.0814
8.9313 | -104.0482
8.9351 | -104.0036
8.9389 | -103.9486
8.9426 | -103.8895
8.9464 | -103.8311
8.9502 | -103.7647
8.954 | -103.6936
8.9578 | -103.6252
8.9615 | -103.5603
8.9653 | -103.5021
8.9691 | -103.4513
8.9729 | -103.4078
8.9767 | -103.3743
8.9804 | -103.352
8.9842 | -103.3399
8.988 | -103.3401
8.9918 | -103.3538
8.9956 | -103.3798
8.9993 | -103.417
9.0031 | -103.4634
9.0069 | -103.5128
9.0107 | -103.5629
9.0145 | -103.6099
9.0182 | -103.6547
9.022 | -103.6996
9.0258 | -103.7408
9.0296 | -103.7765
9.0334 | -103.8003
9.0371 | -103.8253
9.0409 | -103.8516
9.0447 | -103.883
9.0485 | -103.9195
9.0523 | -103.9512
9.056 | -103.9802
9.0598 | -104.0109
9.0636 | -104.0461
9.0674 | -104.0826
9.0712 | -104.1233
9.0749 | -104.1618
9.0787 | -104.1952
9.0825 | -104.2307
9.0863 | -104.2667
9.09 | -104.2937
9.0938 | -104.3125
9.0976 | -104.3248
9.1014 | -104.3305
9.1052 | -104.3244
9.1089 | -104.308
9.1127 | -104.2897
9.1165 | -104.2684
9.1203 | -104.2416
9.1241 | -104.2072
9.1278 | -104.1707
9.1316 | -104.1318
9.1354 | -104.0901
9.1392 | -104.0586
9.143 | -104.0308
9.1467 | -104.0063
9.1505 | -103.9841
9.1543 | -103.968
9.1581 | -103.9554
9.1619 | -103.9489
9.1656 | -103.958
9.1694 | -103.9668
9.1732 | -103.9786
9.177 | -103.9953
9.1808 | -104.0119
9.1845 | -104.0309
9.1883 | -104.0514
9.1921 | -104.0651
9.1959 | -104.0772
9.1997 | -104.0931
9.2034 | -104.1
9.2072 | -104.1012
9.211 | -104.0972
9.2148 | -104.0872
9.2186 | -104.0735
9.2223 | -104.0557
9.2261 | -104.0393
9.2299 | -104.0238
9.2337 | -104.0094
9.2375 | -103.9981
9.2412 | -103.9894
9.245 | -103.9755
9.2488 | -103.9665
9.2526 | -103.9637
9.2564 | -103.969
9.2601 | -103.9766
9.2639 | -103.9894
9.2677 | -104.0138
9.2715 | -104.0321
9.2753 | -104.0496
9.279 | -104.0708
9.2828 | -104.0899
9.2866 | -104.1091
9.2904 | -104.1276
9.2942 | -104.1518
9.2979 | -104.1739
9.3017 | -104.1858
9.3055 | -104.1951
9.3093 | -104.2046
9.313 | -104.212
9.3168 | -104.2159
9.3206 | -104.2202
9.3244 | -104.2225
9.3282 | -104.2243
9.3319 | -104.2174
9.3357 | -104.2037
9.3395 | -104.195
9.3433 | -104.1801
9.3471 | -104.1644
9.3508 | -104.1489
9.3546 | -104.1298
9.3584 | -104.111
9.3622 | -104.0899
9.366 | -104.0688
9.3697 | -104.0506
9.3735 | -104.0368
9.3773 | -104.0227
9.3811 | -104.0146
9.3849 | -104.011
9.3886 | -104.0074
9.3924 | -104.0078
9.3962 | -104.0118
9.4 | -104.0219
9.4038 | -104.0374
9.4075 | -104.0564
9.4113 | -104.0782
9.4151 | -104.1058
9.4189 | -104.1371
9.4227 | -104.1701
9.4264 | -104.2003
9.4302 | -104.2317
9.434 | -104.2609
9.4378 | -104.2887
9.4416 | -104.3199
9.4453 | -104.3545
9.4491 | -104.3787
9.4529 | -104.3926
9.4567 | -104.4086
9.4605 | -104.4121
9.4642 | -104.408
9.468 | -104.4018
9.4718 | -104.3954
9.4756 | -104.3807
9.4794 | -104.3537
9.4831 | -104.3291
9.4869 | -104.3022
9.4907 | -104.2688
9.4945 | -104.235
9.4983 | -104.2048
9.502 | -104.1728
9.5058 | -104.1441
9.5096 | -104.1207
9.5134 | -104.1003
9.5172 | -104.0843
9.5209 | -104.072
9.5247 | -104.0627
9.5285 | -104.0641
9.5323 | -104.0701
9.536 | -104.082
9.5398 | -104.0983
9.5436 | -104.1157
9.5474 | -104.1375
9.5512 | -104.1603
9.5549 | -104.1851
9.5587 | -104.2122
9.5625 | -104.2436
9.5663 | -104.2717
9.5701 | -104.2925
9.5738 | -104.3143
9.5776 | -104.3363
9.5814 | -104.3529
9.5852 | -104.3674
9.589 | -104.3852
9.5927 | -104.3996
9.5965 | -104.4095
9.6003 | -104.4213
9.6041 | -104.4339
9.6079 | -104.4414
9.6116 | -104.4438
9.6154 | -104.4483
9.6192 | -104.4526
9.623 | -104.4541
9.6268 | -104.454
9.6305 | -104.4523
9.6343 | -104.4512
9.6381 | -104.452
9.6419 | -104.4516
9.6457 | -104.4491
9.6494 | -104.4485
9.6532 | -104.4483
9.657 | -104.4489
9.6608 | -104.4518
9.6646 | -104.4567
9.6683 | -104.4567
9.6721 | -104.4545
9.6759 | -104.4538
9.6797 | -104.4492
9.6835 | -104.4467
9.6872 | -104.4377
9.691 | -104.4227
9.6948 | -104.4029
9.6986 | -104.3827
9.7024 | -104.3579
9.7061 | -104.3325
9.7099 | -104.3061
9.7137 | -104.2715
9.7175 | -104.2353
9.7213 | -104.2051
9.725 | -104.1768
9.7288 | -104.151
9.7326 | -104.1326
9.7364 | -104.1159
9.7401 | -104.1033
9.7439 | -104.0954
9.7477 | -104.0938
9.7515 | -104.0961
9.7553 | -104.1038
9.759 | -104.1159
9.7628 | -104.13
9.7666 | -104.1482
9.7704 | -104.1702
9.7742 | -104.1943
9.7779 | -104.2178
9.7817 | -104.2453
9.7855 | -104.2689
9.7893 | -104.2911
9.7931 | -104.3155
9.7968 | -104.3418
9.8006 | -104.3681
9.8044 | -104.3942
9.8082 | -104.4243
9.812 | -104.4626
9.8157 | -104.497
9.8195 | -104.5324
9.8233 | -104.5664
9.8271 | -104.5952
9.8309 | -104.6282
9.8346 | -104.6612
9.8384 | -104.6883
9.8422 | -104.7129
9.846 | -104.7366
9.8498 | -104.7525
9.8535 | -104.7602
9.8573 | -104.7675
9.8611 | -104.7742
9.8649 | -104.7709
9.8687 | -104.7626
9.8724 | -104.7564
9.8762 | -104.7475
9.88 | -104.7355
9.8838 | -104.7225
9.8876 | -104.7084
9.8913 | -104.6946
9.8951 | -104.681
9.8989 | -104.6687
9.9027 | -104.6572
9.9065 | -104.6451
9.9102 | -104.6312
9.914 | -104.6216
9.9178 | -104.6084
9.9216 | -104.5967
9.9254 | -104.5882
9.9291 | -104.5823
9.9329 | -104.5767
9.9367 | -104.5698
9.9405 | -104.5601
9.9443 | -104.5502
9.948 | -104.5426
9.9518 | -104.5363
9.9556 | -104.5302
9.9594 | -104.5232
9.9631 | -104.5216
9.9669 | -104.5195
9.9707 | -104.5218
9.9745 | -104.529
9.9783 | -104.5352
9.982 | -104.5436
9.9858 | -104.5558
9.9896 | -104.5645
9.9934 | -104.572
9.9972 | -104.5858
10.0009 | -104.6048
10.0047 | -104.6233
10.0085 | -104.6333
10.0123 | -104.6423
10.0161 | -104.6522
10.0198 | -104.6555
10.0236 | -104.6587
10.0274 | -104.6579
10.0312 | -104.6563
10.035 | -104.6521
10.0387 | -104.6443
10.0425 | -104.6362
10.0463 | -104.6207
10.0501 | -104.5991
10.0539 | -104.5851
10.0576 | -104.5727
10.0614 | -104.5576
10.0652 | -104.5418
10.069 | -104.5245
10.0728 | -104.5076
10.0765 | -104.4942
10.0803 | -104.4826
10.0841 | -104.4768
10.0879 | -104.4781
10.0917 | -104.4795
10.0954 | -104.4821
10.0992 | -104.4874
10.103 | -104.4935
10.1068 | -104.5052
10.1106 | -104.519
10.1143 | -104.5358
10.1181 | -104.5525
10.1219 | -104.5688
10.1257 | -104.5883
10.1295 | -104.6073
10.1332 | -104.6242
10.137 | -104.6461
10.1408 | -104.6694
10.1446 | -104.6919
10.1484 | -104.7085
10.1521 | -104.7204
10.1559 | -104.737
10.1597 | -104.75
10.1635 | -104.7561
10.1672 | -104.7629
10.171 | -104.763
10.1748 | -104.7638
10.1786 | -104.7587
10.1824 | -104.7493
10.1861 | -104.7392
10.1899 | -104.7277
10.1937 | -104.7174
10.1975 | -104.7084
10.2013 | -104.6894
10.205 | -104.6695
10.2088 | -104.6573
10.2126 | -104.6519
10.2164 | -104.6475
10.2202 | -104.6424
10.2239 | -104.6389
10.2277 | -104.6382
10.2315 | -104.6365
10.2353 | -104.6351
10.2391 | -104.6431
10.2428 | -104.6551
10.2466 | -104.6711
10.2504 | -104.6896
10.2542 | -104.7064
10.258 | -104.7209
10.2617 | -104.7369
10.2655 | -104.7556
10.2693 | -104.7736
10.2731 | -104.7912
10.2769 | -104.8098
10.2806 | -104.825
10.2844 | -104.8373
10.2882 | -104.8493
10.292 | -104.8575
10.2958 | -104.8689
10.2995 | -104.8814
10.3033 | -104.8899
10.3071 | -104.8968
10.3109 | -104.9026
10.3147 | -104.9061
10.3184 | -104.9109
10.3222 | -104.9109
10.326 | -104.9065
10.3298 | -104.8981
10.3336 | -104.8875
10.3373 | -104.8741
10.3411 | -104.862
10.3449 | -104.8488
10.3487 | -104.8361
10.3525 | -104.8235
10.3562 | -104.806
10.36 | -104.797
10.3638 | -104.7851
10.3676 | -104.7749
10.3713 | -104.7661
10.3751 | -104.7572
10.3789 | -104.7501
10.3827 | -104.7489
10.3865 | -104.7504
10.3902 | -104.7501
10.394 | -104.7492
10.3978 | -104.7506
10.4016 | -104.7586
10.4054 | -104.7679
10.4091 | -104.78
10.4129 | -104.7901
10.4167 | -104.8006
10.4205 | -104.8109
10.4243 | -104.8205
10.428 | -104.8296
10.4318 | -104.8367
10.4356 | -104.8412
10.4394 | -104.8435
10.4432 | -104.8441
10.4469 | -104.8428
10.4507 | -104.8422
10.4545 | -104.8406
10.4583 | -104.8399
10.4621 | -104.8399
10.4658 | -104.8366
10.4696 | -104.8303
10.4734 | -104.831
10.4772 | -104.8345
10.481 | -104.839
10.4847 | -104.8392
10.4885 | -104.838
10.4923 | -104.8377
10.4961 | -104.8351
10.4999 | -104.8324
10.5036 | -104.8247
10.5074 | -104.8135
10.5112 | -104.8014
10.515 | -104.7874
10.5188 | -104.7705
10.5225 | -104.7541
10.5263 | -104.7394
10.5301 | -104.7233
10.5339 | -104.7043
10.5377 | -104.685
10.5414 | -104.6658
10.5452 | -104.6444
10.549 | -104.6289
10.5528 | -104.6154
10.5566 | -104.6028
10.5603 | -104.5938
10.5641 | -104.583
10.5679 | -104.5753
10.5717 | -104.5711
10.5755 | -104.5724
10.5792 | -104.5755
10.583 | -104.5812
10.5868 | -104.5923
10.5906 | -104.6057
10.5943 | -104.6231
10.5981 | -104.641
10.6019 | -104.6612
10.6057 | -104.6827
10.6095 | -104.7075
10.6132 | -104.7298
10.617 | -104.7543
10.6208 | -104.7794
10.6246 | -104.8029
10.6284 | -104.8224
10.6321 | -104.8428
10.6359 | -104.8632
10.6397 | -104.8822
10.6435 | -104.8962
10.6473 | -104.9071
10.651 | -104.9213
10.6548 | -104.9373
10.6586 | -104.9529
10.6624 | -104.9671
10.6662 | -104.9804
10.6699 | -104.9946
10.6737 | -105.0096
10.6775 | -105.0264
10.6813 | -105.0387
10.6851 | -105.0536
10.6888 | -105.0722
10.6926 | -105.0932
10.6964 | -105.1157
10.7002 | -105.1409
10.704 | -105.1611
10.7077 | -105.1767
10.7115 | -105.192
10.7153 | -105.2082
10.7191 | -105.2249
10.7229 | -105.2343
10.7266 | -105.242
10.7304 | -105.2456
10.7342 | -105.2452
10.738 | -105.2428
10.7418 | -105.2366
10.7455 | -105.2278
10.7493 | -105.2164
10.7531 | -105.2061
10.7569 | -105.1912
10.7607 | -105.1768
10.7644 | -105.1604
10.7682 | -105.1456
10.772 | -105.1271
10.7758 | -105.1086
10.7796 | -105.0912
10.7833 | -105.0736
10.7871 | -105.0569
10.7909 | -105.045
10.7947 | -105.0374
10.7985 | -105.0306
10.8022 | -105.0175
10.806 | -105.0073
10.8098 | -104.9985
10.8136 | -104.9926
10.8173 | -104.9901
10.8211 | -104.9898
10.8249 | -104.988
10.8287 | -104.9915
10.8325 | -104.9909
10.8362 | -104.9969
10.84 | -105.0048
10.8438 | -105.0097
10.8476 | -105.0113
10.8514 | -105.014
10.8551 | -105.0223
10.8589 | -105.0289
10.8627 | -105.0361
10.8665 | -105.0409
10.8703 | -105.0426
10.874 | -105.041
10.8778 | -105.0386
10.8816 | -105.0312
10.8854 | -105.0223
10.8892 | -105.0103
10.8929 | -104.9963
10.8967 | -104.9822
10.9005 | -104.9656
10.9043 | -104.9532
10.9081 | -104.9419
10.9118 | -104.934
10.9156 | -104.9278
10.9194 | -104.9202
10.9232 | -104.9158
10.927 | -104.9171
10.9307 | -104.9195
10.9345 | -104.9197
10.9383 | -104.9209
10.9421 | -104.9252
10.9459 | -104.9336
10.9496 | -104.9428
10.9534 | -104.9513
10.9572 | -104.963
10.961 | -104.9809
10.9648 | -104.9962
10.9685 | -105.0086
10.9723 | -105.0223
10.9761 | -105.0368
10.9799 | -105.0499
10.9837 | -105.0618
10.9874 | -105.0689
10.9912 | -105.0783
10.995 | -105.0824
10.9988 | -105.0847
11.0026 | -105.0792
11.0063 | -105.0707
11.0101 | -105.0636
11.0139 | -105.0536
11.0177 | -105.045
11.0214 | -105.0332
11.0252 | -105.0211
11.029 | -105.0089
11.0328 | -104.9982
11.0366 | -104.9862
11.0403 | -104.9729
11.0441 | -104.9628
11.0479 | -104.9537
11.0517 | -104.9452
11.0555 | -104.9411
11.0592 | -104.937
11.063 | -104.9296
11.0668 | -104.9202
11.0706 | -104.9125
11.0744 | -104.9062
11.0781 | -104.9037
11.0819 | -104.9012
11.0857 | -104.8971
11.0895 | -104.8929
11.0933 | -104.8887
11.097 | -104.8902
11.1008 | -104.8888
11.1046 | -104.8833
11.1084 | -104.8791
11.1122 | -104.8744
11.1159 | -104.8712
11.1197 | -104.87
11.1235 | -104.8671
11.1273 | -104.8652
11.1311 | -104.8649
11.1348 | -104.8672
11.1386 | -104.8693
11.1424 | -104.8736
11.1462 | -104.8797
11.15 | -104.8856
11.1537 | -104.8948
11.1575 | -104.9035
11.1613 | -104.9119
11.1651 | -104.929
11.1689 | -104.9478
11.1726 | -104.9629
11.1764 | -104.973
11.1802 | -104.9853
11.184 | -105.0015
11.1878 | -105.0173
11.1915 | -105.036
11.1953 | -105.0537
11.1991 | -105.0711
11.2029 | -105.0854
11.2067 | -105.0946
11.2104 | -105.1068
11.2142 | -105.12
11.218 | -105.1253
11.2218 | -105.132
11.2256 | -105.141
11.2293 | -105.1476
11.2331 | -105.1481
11.2369 | -105.1472
11.2407 | -105.1454
11.2444 | -105.1414
11.2482 | -105.1404
11.252 | -105.1391
11.2558 | -105.1359
11.2596 | -105.1288
11.2633 | -105.1238
11.2671 | -105.1164
11.2709 | -105.1088
11.2747 | -105.099
11.2785 | -105.0893
11.2822 | -105.0797
11.286 | -105.0739
11.2898 | -105.0674
11.2936 | -105.0611
11.2974 | -105.0507
11.3011 | -105.0452
11.3049 | -105.0428
11.3087 | -105.0431
11.3125 | -105.0444
11.3163 | -105.0452
11.32 | -105.049
11.3238 | -105.0532
11.3276 | -105.0526
11.3314 | -105.0557
11.3352 | -105.0611
11.3389 | -105.0673
11.3427 | -105.0721
11.3465 | -105.0747
11.3503 | -105.0779
11.3541 | -105.0785
11.3578 | -105.0795
11.3616 | -105.0774
11.3654 | -105.0754
11.3692 | -105.0767
11.373 | -105.077
11.3767 | -105.0687
11.3805 | -105.0608
11.3843 | -105.0501
11.3881 | -105.0405
11.3919 | -105.0343
11.3956 | -105.026
11.3994 | -105.0133
11.4032 | -105.0007
11.407 | -104.9894
11.4108 | -104.9752
11.4145 | -104.9604
11.4183 | -104.9463
11.4221 | -104.936
11.4259 | -104.9266
11.4297 | -104.9171
11.4334 | -104.9061
11.4372 | -104.8999
11.441 | -104.8963
11.4448 | -104.8949
11.4485 | -104.8971
11.4523 | -104.894
11.4561 | -104.8969
11.4599 | -104.9023
11.4637 | -104.9106
11.4674 | -104.9181
11.4712 | -104.9247
11.475 | -104.9338
11.4788 | -104.9435
11.4826 | -104.9544
11.4863 | -104.9677
11.4901 | -104.9826
11.4939 | -104.9997
11.4977 | -105.0093
11.5015 | -105.0206
11.5052 | -105.0369
11.509 | -105.0525
11.5128 | -105.066
11.5166 | -105.081
11.5204 | -105.0961
11.5241 | -105.112
11.5279 | -105.1262
11.5317 | -105.134
11.5355 | -105.137
11.5393 | -105.1394
11.543 | -105.1457
11.5468 | -105.1493
11.5506 | -105.1486
11.5544 | -105.1402
11.5582 | -105.1313
11.5619 | -105.1243
11.5657 | -105.1115
11.5695 | -105.0944
11.5733 | -105.0806
11.5771 | -105.0661
11.5808 | -105.0476
11.5846 | -105.0327
11.5884 | -105.0152
11.5922 | -104.9998
11.596 | -104.9893
11.5997 | -104.9791
11.6035 | -104.9713
11.6073 | -104.9644
11.6111 | -104.9596
11.6149 | -104.9579
11.6186 | -104.9624
11.6224 | -104.9678
11.6262 | -104.9745
11.63 | -104.9859
11.6338 | -105.0006
11.6375 | -105.0183
11.6413 | -105.0386
11.6451 | -105.0646
11.6489 | -105.0908
11.6526 | -105.114
11.6564 | -105.1383
11.6602 | -105.1604
11.664 | -105.183
11.6678 | -105.2078
11.6715 | -105.2241
11.6753 | -105.238
11.6791 | -105.2507
11.6829 | -105.2605
11.6867 | -105.2635
11.6904 | -105.2638
11.6942 | -105.2592
11.698 | -105.2532
11.7018 | -105.2514
11.7056 | -105.2478
11.7093 | -105.2389
11.7131 | -105.227
11.7169 | -105.2157
11.7207 | -105.2044
11.7245 | -105.1936
11.7282 | -105.1816
11.732 | -105.1702
11.7358 | -105.1615
11.7396 | -105.1569
11.7434 | -105.1524
11.7471 | -105.149
11.7509 | -105.151
11.7547 | -105.1541
11.7585 | -105.1605
11.7623 | -105.1684
11.766 | -105.1801
11.7698 | -105.1953
11.7736 | -105.2104
11.7774 | -105.2275
11.7812 | -105.2469
11.7849 | -105.2732
11.7887 | -105.3004
11.7925 | -105.3257
11.7963 | -105.3487
11.8001 | -105.3725
11.8038 | -105.3981
11.8076 | -105.4217
11.8114 | -105.4429
11.8152 | -105.4598
11.819 | -105.4793
11.8227 | -105.4967
11.8265 | -105.5061
11.8303 | -105.5157
11.8341 | -105.518
11.8379 | -105.5179
11.8416 | -105.5161
11.8454 | -105.5106
11.8492 | -105.5018
11.853 | -105.4939
11.8568 | -105.4849
11.8605 | -105.4728
11.8643 | -105.4591
11.8681 | -105.4455
11.8719 | -105.4248
11.8757 | -105.402
11.8794 | -105.3794
11.8832 | -105.3612
11.887 | -105.3445
11.8908 | -105.3281
11.8945 | -105.3139
11.8983 | -105.3003
11.9021 | -105.2902
11.9059 | -105.282
11.9097 | -105.2697
11.9134 | -105.259
11.9172 | -105.2527
11.921 | -105.2483
11.9248 | -105.2484
11.9286 | -105.2501
11.9323 | -105.253
11.9361 | -105.2574
11.9399 | -105.2673
11.9437 | -105.2789
11.9475 | -105.2903
11.9512 | -105.2995
11.955 | -105.3078
11.9588 | -105.3185
11.9626 | -105.3313
11.9664 | -105.3467
11.9701 | -105.3642
11.9739 | -105.3778
11.9777 | -105.3882
11.9815 | -105.4041
11.9853 | -105.417
11.989 | -105.4266
11.9928 | -105.4355
11.9966 | -105.4471
12.0004 | -105.4577
12.0042 | -105.4664
12.0079 | -105.4777
12.0117 | -105.485
12.0155 | -105.4927
12.0193 | -105.4979
12.0231 | -105.5052
12.0268 | -105.512
12.0306 | -105.5202
12.0344 | -105.5292
12.0382 | -105.5381
12.042 | -105.5469
12.0457 | -105.5597
12.0495 | -105.5716
12.0533 | -105.5857
12.0571 | -105.5965
12.0609 | -105.6037
12.0646 | -105.6118
12.0684 | -105.6162
12.0722 | -105.6191
12.076 | -105.621
12.0798 | -105.6224
12.0835 | -105.6248
12.0873 | -105.6174
12.0911 | -105.6069
12.0949 | -105.5947
12.0986 | -105.5781
12.1024 | -105.5617
12.1062 | -105.5455
12.11 | -105.5262
12.1138 | -105.5054
12.1175 | -105.4853
12.1213 | -105.4668
12.1251 | -105.4504
12.1289 | -105.4345
12.1327 | -105.4206
12.1364 | -105.4094
12.1402 | -105.4004
12.144 | -105.3896
12.1478 | -105.3852
12.1516 | -105.387
12.1553 | -105.3891
12.1591 | -105.3971
12.1629 | -105.4097
12.1667 | -105.4229
12.1705 | -105.4385
12.1742 | -105.4566
12.178 | -105.4757
12.1818 | -105.492
12.1856 | -105.5065
12.1894 | -105.5199
12.1931 | -105.5311
12.1969 | -105.5404
12.2007 | -105.5458
12.2045 | -105.5512
12.2083 | -105.5519
12.212 | -105.5504
12.2158 | -105.5477
12.2196 | -105.5429
12.2234 | -105.5305
12.2272 | -105.5162
12.2309 | -105.4989
12.2347 | -105.4797
12.2385 | -105.4622
12.2423 | -105.4432
12.2461 | -105.4245
12.2498 | -105.4058
12.2536 | -105.3868
12.2574 | -105.3661
12.2612 | -105.3466
12.265 | -105.3354
12.2687 | -105.3219
12.2725 | -105.3075
12.2763 | -105.2939
12.2801 | -105.2799
12.2839 | -105.2698
12.2876 | -105.2651
12.2914 | -105.263
12.2952 | -105.2599
12.299 | -105.2591
12.3027 | -105.2615
12.3065 | -105.2648
12.3103 | -105.2681
12.3141 | -105.2753
12.3179 | -105.2828
12.3216 | -105.2914
12.3254 | -105.3015
12.3292 | -105.3102
12.333 | -105.3179
12.3368 | -105.327
12.3405 | -105.3343
12.3443 | -105.3435
12.3481 | -105.353
12.3519 | -105.3614
12.3557 | -105.3688
12.3594 | -105.3773
12.3632 | -105.386
12.367 | -105.3948
12.3708 | -105.4027
12.3746 | -105.4108
12.3783 | -105.4205
12.3821 | -105.4296
12.3859 | -105.4378
12.3897 | -105.4439
12.3935 | -105.4516
12.3972 | -105.4587
12.401 | -105.4666
12.4048 | -105.4743
12.4086 | -105.4826
12.4124 | -105.4873
12.4161 | -105.4833
12.4199 | -105.4825
12.4237 | -105.4856
12.4275 | -105.4851
12.4313 | -105.4844
12.435 | -105.4875
12.4388 | -105.4846
12.4426 | -105.479
12.4464 | -105.476
12.4502 | -105.4708
12.4539 | -105.4677
12.4577 | -105.4656
12.4615 | -105.4584
12.4653 | -105.4511
12.4691 | -105.445
12.4728 | -105.4404
12.4766 | -105.4356
12.4804 | -105.4272
12.4842 | -105.4214
12.488 | -105.4155
12.4917 | -105.4103
12.4955 | -105.4105
12.4993 | -105.4146
12.5031 | -105.4196
12.5069 | -105.4257
12.5106 | -105.4353
12.5144 | -105.443
12.5182 | -105.455
12.522 | -105.4652
12.5257 | -105.4749
12.5295 | -105.4856
12.5333 | -105.494
12.5371 | -105.5006
12.5409 | -105.5042
12.5446 | -105.5087
12.5484 | -105.51
12.5522 | -105.5119
12.556 | -105.511
12.5598 | -105.5095
12.5635 | -105.5095
12.5673 | -105.5089
12.5711 | -105.5102
12.5749 | -105.5086
12.5787 | -105.5067
12.5824 | -105.5054
12.5862 | -105.5051
12.59 | -105.5049
12.5938 | -105.5078
12.5976 | -105.5117
12.6013 | -105.505
12.6051 | -105.4964
12.6089 | -105.4891
12.6127 | -105.482
12.6165 | -105.4737
12.6202 | -105.4695
12.624 | -105.4645
12.6278 | -105.4569
12.6316 | -105.4505
12.6354 | -105.445
12.6391 | -105.4406
12.6429 | -105.439
12.6467 | -105.4328
12.6505 | -105.4268
12.6543 | -105.4249
12.658 | -105.4211
12.6618 | -105.4192
12.6656 | -105.4186
12.6694 | -105.4144
12.6732 | -105.4102
12.6769 | -105.4035
12.6807 | -105.3931
12.6845 | -105.3854
12.6883 | -105.3801
12.6921 | -105.3722
12.6958 | -105.3656
12.6996 | -105.3559
12.7034 | -105.3481
12.7072 | -105.342
12.711 | -105.3382
12.7147 | -105.3352
12.7185 | -105.3301
12.7223 | -105.3216
12.7261 | -105.3106
12.7299 | -105.3032
12.7336 | -105.2985
12.7374 | -105.2932
12.7412 | -105.2895
12.745 | -105.2857
12.7487 | -105.2814
12.7525 | -105.2786
12.7563 | -105.2743
12.7601 | -105.2751
12.7639 | -105.2814
12.7676 | -105.2901
12.7714 | -105.2979
12.7752 | -105.3066
12.779 | -105.319
12.7828 | -105.337
12.7865 | -105.358
12.7903 | -105.3766
12.7941 | -105.396
12.7979 | -105.4153
12.8017 | -105.4357
12.8054 | -105.4551
12.8092 | -105.4735
12.813 | -105.4952
12.8168 | -105.5157
12.8206 | -105.5317
12.8243 | -105.5469
12.8281 | -105.5633
12.8319 | -105.5791
12.8357 | -105.5953
12.8395 | -105.6084
12.8432 | -105.6206
12.847 | -105.6295
12.8508 | -105.6358
12.8546 | -105.6492
12.8584 | -105.6615
12.8621 | -105.6737
12.8659 | -105.6905
12.8697 | -105.7025
12.8735 | -105.7142
12.8773 | -105.7286
12.881 | -105.7389
12.8848 | -105.755
12.8886 | -105.7723
12.8924 | -105.7895
12.8962 | -105.8041
12.8999 | -105.8227
12.9037 | -105.8415
12.9075 | -105.8607
12.9113 | -105.8795
12.9151 | -105.8955
12.9188 | -105.9078
12.9226 | -105.9161
12.9264 | -105.9227
12.9302 | -105.9279
12.934 | -105.9284
12.9377 | -105.9269
12.9415 | -105.9272
12.9453 | -105.9197
12.9491 | -105.9045
12.9528 | -105.885
12.9566 | -105.8641
12.9604 | -105.8425
12.9642 | -105.8176
12.968 | -105.7895
12.9717 | -105.7635
12.9755 | -105.7372
12.9793 | -105.7066
12.9831 | -105.6779
12.9869 | -105.6537
12.9906 | -105.6311
12.9944 | -105.61
12.9982 | -105.5906
13.002 | -105.5759
13.0058 | -105.5643
13.0095 | -105.5572
13.0133 | -105.5534
13.0171 | -105.5587
13.0209 | -105.5624
13.0247 | -105.5729
13.0284 | -105.5894
13.0322 | -105.6089
13.036 | -105.6285
13.0398 | -105.6453
13.0436 | -105.6649
13.0473 | -105.686
13.0511 | -105.7033
13.0549 | -105.723
13.0587 | -105.7425
13.0625 | -105.7607
13.0662 | -105.7778
13.07 | -105.7856
13.0738 | -105.7876
13.0776 | -105.7872
13.0814 | -105.7876
13.0851 | -105.7663
13.0889 | -105.7587
13.0927 | -105.7493
13.0965 | -105.7372
13.1003 | -105.7232
13.104 | -105.7093
13.1078 | -105.6958
13.1116 | -105.6819
13.1154 | -105.667
13.1192 | -105.652
13.1229 | -105.6418
13.1267 | -105.6343
13.1305 | -105.6248
13.1343 | -105.6149
13.1381 | -105.6072
13.1418 | -105.6013
13.1456 | -105.5996
13.1494 | -105.5982
13.1532 | -105.5939
13.157 | -105.592
13.1607 | -105.5911
13.1645 | -105.5884
13.1683 | -105.5901
13.1721 | -105.5884
13.1758 | -105.5857
13.1796 | -105.5813
13.1834 | -105.5765
13.1872 | -105.5727
13.191 | -105.5653
13.1947 | -105.5583
13.1985 | -105.551
13.2023 | -105.5419
13.2061 | -105.5305
13.2099 | -105.5206
13.2136 | -105.5031
13.2174 | -105.4858
13.2212 | -105.4706
13.225 | -105.4554
13.2288 | -105.4417
13.2325 | -105.4278
13.2363 | -105.4097
13.2401 | -105.3968
13.2439 | -105.3854
13.2477 | -105.3744
13.2514 | -105.3656
13.2552 | -105.361
13.259 | -105.3594
13.2628 | -105.3573
13.2666 | -105.356
13.2703 | -105.357
13.2741 | -105.3645
13.2779 | -105.3725
13.2817 | -105.3831
13.2855 | -105.3937
13.2892 | -105.4058
13.293 | -105.4205
13.2968 | -105.4345
13.3006 | -105.4457
13.3044 | -105.4591
13.3081 | -105.472
13.3119 | -105.4854
13.3157 | -105.4991
13.3195 | -105.5116
13.3233 | -105.5236
13.327 | -105.5371
13.3308 | -105.5525
13.3346 | -105.5692
13.3384 | -105.5823
13.3422 | -105.5928
13.3459 | -105.6062
13.3497 | -105.6181
13.3535 | -105.6294
13.3573 | -105.6451
13.3611 | -105.6569
13.3648 | -105.6674
13.3686 | -105.676
13.3724 | -105.6881
13.3762 | -105.7017
13.3799 | -105.7114
13.3837 | -105.7251
13.3875 | -105.7383
13.3913 | -105.7496
13.3951 | -105.7604
13.3988 | -105.7711
13.4026 | -105.7821
13.4064 | -105.7945
13.4102 | -105.8034
13.414 | -105.8086
13.4177 | -105.8139
13.4215 | -105.8179
13.4253 | -105.8192
13.4291 | -105.8202
13.4329 | -105.8188
13.4366 | -105.8138
13.4404 | -105.8066
13.4442 | -105.7998
13.448 | -105.7915
13.4518 | -105.7831
13.4555 | -105.7778
13.4593 | -105.7719
13.4631 | -105.7627
13.4669 | -105.7516
13.4707 | -105.7456
13.4744 | -105.7402
13.4782 | -105.737
13.482 | -105.7381
13.4858 | -105.7354
13.4896 | -105.7304
13.4933 | -105.7304
13.4971 | -105.7322
13.5009 | -105.7354
13.5047 | -105.7379
13.5085 | -105.7394
13.5122 | -105.7414
13.516 | -105.7459
13.5198 | -105.7485
13.5236 | -105.7503
13.5274 | -105.7556
13.5311 | -105.7614
13.5349 | -105.7682
13.5387 | -105.7755
13.5425 | -105.7826
13.5463 | -105.7898
13.55 | -105.7969
13.5538 | -105.8072
13.5576 | -105.8185
13.5614 | -105.8273
13.5652 | -105.8351
13.5689 | -105.8415
13.5727 | -105.8506
13.5765 | -105.8577
13.5803 | -105.8668
13.584 | -105.8766
13.5878 | -105.8848
13.5916 | -105.8889
13.5954 | -105.8909
13.5992 | -105.8934
13.6029 | -105.8926
13.6067 | -105.8894
13.6105 | -105.8842
13.6143 | -105.8788
13.6181 | -105.8699
13.6218 | -105.859
13.6256 | -105.8417
13.6294 | -105.8136
13.6332 | -105.7936
13.637 | -105.7735
13.6407 | -105.752
13.6445 | -105.7325
13.6483 | -105.7124
13.6521 | -105.6925
13.6559 | -105.6729
13.6596 | -105.6548
13.6634 | -105.6373
13.6672 | -105.6206
13.671 | -105.6055
13.6748 | -105.5942
13.6785 | -105.5835
13.6823 | -105.58
13.6861 | -105.5807
13.6899 | -105.5792
13.6937 | -105.5788
13.6974 | -105.5794
13.7012 | -105.5849
13.705 | -105.5934
13.7088 | -105.6056
13.7126 | -105.6193
13.7163 | -105.6374
13.7201 | -105.6575
13.7239 | -105.678
13.7277 | -105.7028
13.7315 | -105.7294
13.7352 | -105.7569
13.739 | -105.7872
13.7428 | -105.8175
13.7466 | -105.8478
13.7504 | -105.8776
13.7541 | -105.9103
13.7579 | -105.9428
13.7617 | -105.974
13.7655 | -106.0066
13.7693 | -106.0348
13.773 | -106.0607
13.7768 | -106.0831
13.7806 | -106.1031
13.7844 | -106.1237
13.7882 | -106.1455
13.7919 | -106.1663
13.7957 | -106.1824
13.7995 | -106.1957
13.8033 | -106.2082
13.807 | -106.2163
13.8108 | -106.2191
13.8146 | -106.221
13.8184 | -106.2228
13.8222 | -106.2251
13.8259 | -106.2241
13.8297 | -106.2203
13.8335 | -106.2163
13.8373 | -106.2081
13.8411 | -106.201
13.8448 | -106.1908
13.8486 | -106.1782
13.8524 | -106.1659
13.8562 | -106.1541
13.86 | -106.1412
13.8637 | -106.1298
13.8675 | -106.1167
13.8713 | -106.1016
13.8751 | -106.0857
13.8789 | -106.068
13.8826 | -106.0491
13.8864 | -106.033
13.8902 | -106.0153
13.894 | -105.9944
13.8978 | -105.9762
13.9015 | -105.9539
13.9053 | -105.9329
13.9091 | -105.9099
13.9129 | -105.8864
13.9167 | -105.863
13.9204 | -105.8409
13.9242 | -105.8183
13.928 | -105.7956
13.9318 | -105.7771
13.9356 | -105.7623
13.9393 | -105.7426
13.9431 | -105.7245
13.9469 | -105.7108
13.9507 | -105.6992
13.9545 | -105.6866
13.9582 | -105.6784
13.962 | -105.6696
13.9658 | -105.6584
13.9696 | -105.6486
13.9734 | -105.6428
13.9771 | -105.6378
13.9809 | -105.6336
13.9847 | -105.6293
13.9885 | -105.6247
13.9923 | -105.6216
13.996 | -105.6212
13.9998 | -105.6215
14.0036 | -105.6216
14.0074 | -105.6243
14.0112 | -105.6271
14.0149 | -105.6335
14.0187 | -105.6417
14.0225 | -105.6494
14.0263 | -105.6594
14.03 | -105.6709
14.0338 | -105.6831
14.0376 | -105.6959
14.0414 | -105.7067
14.0452 | -105.7229
14.0489 | -105.7443
14.0527 | -105.7621
14.0565 | -105.7793
14.0603 | -105.7971
14.0641 | -105.8136
14.0678 | -105.8304
14.0716 | -105.8493
14.0754 | -105.8697
14.0792 | -105.8897
14.083 | -105.9097
14.0867 | -105.9237
14.0905 | -105.9405
14.0943 | -105.9558
14.0981 | -105.9666
14.1019 | -105.9759
14.1056 | -105.9823
14.1094 | -105.9903
14.1132 | -105.9958
14.117 | -105.9972
14.1208 | -105.9959
14.1245 | -105.9952
14.1283 | -105.9937
14.1321 | -105.9866
14.1359 | -105.9761
14.1397 | -105.9708
14.1434 | -105.9655
14.1472 | -105.956
14.151 | -105.9487
14.1548 | -105.9417
14.1586 | -105.935
14.1623 | -105.9319
14.1661 | -105.931
14.1699 | -105.9299
14.1737 | -105.9402
14.1775 | -105.9428
14.1812 | -105.948
14.185 | -105.954
14.1888 | -105.9607
14.1926 | -105.9695
14.1964 | -105.9786
14.2001 | -105.9893
14.2039 | -106.0027
14.2077 | -106.0164
14.2115 | -106.0316
14.2153 | -106.0466
14.219 | -106.0609
14.2228 | -106.0772
14.2266 | -106.0934
14.2304 | -106.107
14.2341 | -106.1174
14.2379 | -106.1268
14.2417 | -106.1377
14.2455 | -106.1478
14.2493 | -106.1545
14.253 | -106.1567
14.2568 | -106.1618
14.2606 | -106.1675
14.2644 | -106.1697
14.2682 | -106.169
14.2719 | -106.1668
14.2757 | -106.1633
14.2795 | -106.1564
14.2833 | -106.1522
14.2871 | -106.1477
14.2908 | -106.1423
14.2946 | -106.1314
14.2984 | -106.1225
14.3022 | -106.1145
14.306 | -106.1061
14.3097 | -106.0986
14.3135 | -106.0898
14.3173 | -106.0814
14.3211 | -106.073
14.3249 | -106.0665
14.3286 | -106.0583
14.3324 | -106.0494
14.3362 | -106.0438
14.34 | -106.0359
14.3438 | -106.0296
14.3475 | -106.0253
14.3513 | -106.0235
14.3551 | -106.0196
14.3589 | -106.0178
14.3627 | -106.0155
14.3664 | -106.0123
14.3702 | -106.0123
14.374 | -106.0164
14.3778 | -106.0207
14.3816 | -106.0237
14.3853 | -106.0278
14.3891 | -106.0339
14.3929 | -106.0425
14.3967 | -106.0551
14.4005 | -106.0676
14.4042 | -106.0796
14.408 | -106.0914
14.4118 | -106.1024
14.4156 | -106.1162
14.4194 | -106.1313
14.4231 | -106.1496
14.4269 | -106.1655
14.4307 | -106.1773
14.4345 | -106.1914
14.4383 | -106.2077
14.442 | -106.2253
14.4458 | -106.2399
14.4496 | -106.2513
14.4534 | -106.2595
14.4571 | -106.2672
14.4609 | -106.2748
14.4647 | -106.2778
14.4685 | -106.2786
14.4723 | -106.2803
14.476 | -106.2805
14.4798 | -106.2769
14.4836 | -106.2721
14.4874 | -106.2672
14.4912 | -106.2576
14.4949 | -106.2452
14.4987 | -106.2345
14.5025 | -106.2287
14.5063 | -106.22
14.5101 | -106.2064
14.5138 | -106.1938
14.5176 | -106.1806
14.5214 | -106.1673
14.5252 | -106.1531
14.529 | -106.1418
14.5327 | -106.1264
14.5365 | -106.1138
14.5403 | -106.1033
14.5441 | -106.092
14.5479 | -106.0827
14.5516 | -106.0724
14.5554 | -106.0612
14.5592 | -106.049
14.563 | -106.0375
14.5668 | -106.0255
14.5705 | -106.0157
14.5743 | -106.0075
14.5781 | -105.9976
14.5819 | -105.9864
14.5857 | -105.9732
14.5894 | -105.9646
14.5932 | -105.9566
14.597 | -105.9502
14.6008 | -105.944
14.6046 | -105.9363
14.6083 | -105.9332
14.6121 | -105.9327
14.6159 | -105.9329
14.6197 | -105.9353
14.6235 | -105.938
14.6272 | -105.9434
14.631 | -105.9464
14.6348 | -105.9514
14.6386 | -105.9596
14.6424 | -105.9695
14.6461 | -105.9791
14.6499 | -105.9901
14.6537 | -106.0024
14.6575 | -106.0151
14.6612 | -106.0287
14.665 | -106.0398
14.6688 | -106.0489
14.6726 | -106.0601
14.6764 | -106.0693
14.6801 | -106.0771
14.6839 | -106.0823
14.6877 | -106.0871
14.6915 | -106.0914
14.6953 | -106.0953
14.699 | -106.098
14.7028 | -106.0985
14.7066 | -106.0995
14.7104 | -106.0947
14.7142 | -106.0887
14.7179 | -106.0889
14.7217 | -106.0864
14.7255 | -106.0862
14.7293 | -106.0839
14.7331 | -106.0786
14.7368 | -106.0755
14.7406 | -106.073
14.7444 | -106.0711
14.7482 | -106.0676
14.752 | -106.0621
14.7557 | -106.0572
14.7595 | -106.0512
14.7633 | -106.044
14.7671 | -106.0396
14.7709 | -106.0356
14.7746 | -106.0308
14.7784 | -106.0268
14.7822 | -106.0189
14.786 | -106.0127
14.7898 | -106.0083
14.7935 | -106.0046
14.7973 | -106.0015
14.8011 | -105.9959
14.8049 | -105.9932
14.8087 | -105.987
14.8124 | -105.982
14.8162 | -105.9801
14.82 | -105.9782
14.8238 | -105.9792
14.8276 | -105.9824
14.8313 | -105.9866
14.8351 | -105.9899
14.8389 | -105.9925
14.8427 | -106.0004
14.8465 | -106.0102
14.8502 | -106.0208
14.854 | -106.0306
14.8578 | -106.0421
14.8616 | -106.051
14.8653 | -106.0599
14.8691 | -106.0723
14.8729 | -106.0847
14.8767 | -106.0965
14.8805 | -106.1101
14.8842 | -106.1236
14.888 | -106.1376
14.8918 | -106.1498
14.8956 | -106.1612
14.8994 | -106.1708
14.9031 | -106.178
14.9069 | -106.1875
14.9107 | -106.1979
14.9145 | -106.209
14.9183 | -106.2161
14.922 | -106.2225
14.9258 | -106.2276
14.9296 | -106.2292
14.9334 | -106.2285
14.9372 | -106.2292
14.9409 | -106.2285
14.9447 | -106.2281
14.9485 | -106.2287
14.9523 | -106.2266
14.9561 | -106.2224
14.9598 | -106.2166
14.9636 | -106.2097
14.9674 | -106.2049
14.9712 | -106.1998
14.975 | -106.1949
14.9787 | -106.1872
14.9825 | -106.1781
14.9863 | -106.1706
14.9901 | -106.1659
14.9939 | -106.164
14.9976 | -106.1638
15.0014 | -106.1624
15.0052 | -106.1601
15.009 | -106.1608
15.0128 | -106.1652
15.0165 | -106.1706
15.0203 | -106.173
15.0241 | -106.1795
15.0279 | -106.1875
15.0317 | -106.1951
15.0354 | -106.2002
15.0392 | -106.2026
15.043 | -106.2065
15.0468 | -106.2115
15.0506 | -106.2165
15.0543 | -106.2215
15.0581 | -106.2272
15.0619 | -106.2289
15.0657 | -106.229
15.0695 | -106.2295
15.0732 | -106.2291
15.077 | -106.2321
15.0808 | -106.2341
15.0846 | -106.232
15.0883 | -106.2292
15.0921 | -106.2241
15.0959 | -106.2158
15.0997 | -106.2067
15.1035 | -106.1948
15.1072 | -106.1797
15.111 | -106.1692
15.1148 | -106.1562
15.1186 | -106.1401
15.1224 | -106.1221
15.1261 | -106.1024
15.1299 | -106.0816
15.1337 | -106.0597
15.1375 | -106.0347
15.1413 | -106.0091
15.145 | -105.984
15.1488 | -105.9588
15.1526 | -105.932
15.1564 | -105.9053
15.1602 | -105.8809
15.1639 | -105.8611
15.1677 | -105.8411
15.1715 | -105.8176
15.1753 | -105.7947
15.1791 | -105.7765
15.1828 | -105.7561
15.1866 | -105.7374
15.1904 | -105.7212
15.1942 | -105.7093
15.198 | -105.6975
15.2017 | -105.6878
15.2055 | -105.6814
15.2093 | -105.6742
15.2131 | -105.6709
15.2169 | -105.6728
15.2206 | -105.6735
15.2244 | -105.674
15.2282 | -105.6773
15.232 | -105.6839
15.2358 | -105.6921
15.2395 | -105.6992
15.2433 | -105.7066
15.2471 | -105.7176
15.2509 | -105.7301
15.2547 | -105.7437
15.2584 | -105.7588
15.2622 | -105.7745
15.266 | -105.7909
15.2698 | -105.8061
15.2736 | -105.8204
15.2773 | -105.8312
15.2811 | -105.8472
15.2849 | -105.8639
15.2887 | -105.8787
15.2925 | -105.8931
15.2962 | -105.9067
15.3 | -105.9195
15.3038 | -105.9323
15.3076 | -105.945
15.3113 | -105.9592
15.3151 | -105.9713
15.3189 | -105.9808
15.3227 | -105.9892
15.3265 | -106.0006
15.3302 | -106.01
15.334 | -106.0178
15.3378 | -106.0249
15.3416 | -106.0299
15.3454 | -106.0337
15.3491 | -106.0348
15.3529 | -106.0348
15.3567 | -106.0382
15.3605 | -106.0377
15.3643 | -106.0291
15.368 | -106.0246
15.3718 | -106.0169
15.3756 | -106.0073
15.3794 | -105.9977
15.3832 | -105.9851
15.3869 | -105.9724
15.3907 | -105.9629
15.3945 | -105.9542
15.3983 | -105.9407
15.4021 | -105.9295
15.4058 | -105.9188
15.4096 | -105.9055
15.4134 | -105.8918
15.4172 | -105.8781
15.421 | -105.8673
15.4247 | -105.8581
15.4285 | -105.8511
15.4323 | -105.844
15.4361 | -105.8349
15.4399 | -105.8321
15.4436 | -105.8317
15.4474 | -105.8304
15.4512 | -105.8306
15.455 | -105.8321
15.4588 | -105.8361
15.4625 | -105.8395
15.4663 | -105.8402
15.4701 | -105.8447
15.4739 | -105.8513
15.4777 | -105.8539
15.4814 | -105.8591
15.4852 | -105.8673
15.489 | -105.877
15.4928 | -105.8866
15.4966 | -105.896
15.5003 | -105.9058
15.5041 | -105.9124
15.5079 | -105.9182
15.5117 | -105.9247
15.5154 | -105.9315
15.5192 | -105.9346
15.523 | -105.9399
15.5268 | -105.9453
15.5306 | -105.9491
15.5343 | -105.9527
15.5381 | -105.9562
15.5419 | -105.9553
15.5457 | -105.9552
15.5495 | -105.9555
15.5532 | -105.9501
15.557 | -105.9443
15.5608 | -105.9397
15.5646 | -105.9353
15.5684 | -105.9334
15.5721 | -105.933
15.5759 | -105.9314
15.5797 | -105.9253
15.5835 | -105.921
15.5873 | -105.9213
15.591 | -105.9178
15.5948 | -105.9162
15.5986 | -105.9159
15.6024 | -105.9121
15.6062 | -105.9109
15.6099 | -105.9141
15.6137 | -105.9192
15.6175 | -105.9247
15.6213 | -105.9302
15.6251 | -105.9372
15.6288 | -105.9466
15.6326 | -105.9595
15.6364 | -105.9734
15.6402 | -105.9895
15.644 | -106.0039
15.6477 | -106.0171
15.6515 | -106.0311
15.6553 | -106.0467
15.6591 | -106.0664
15.6629 | -106.086
15.6666 | -106.1061
15.6704 | -106.1257
15.6742 | -106.1459
15.678 | -106.1664
15.6818 | -106.1852
15.6855 | -106.2072
15.6893 | -106.2267
15.6931 | -106.2395
15.6969 | -106.2486
15.7007 | -106.2602
15.7044 | -106.2678
15.7082 | -106.2742
15.712 | -106.2765
15.7158 | -106.2753
15.7195 | -106.2729
15.7233 | -106.2697
15.7271 | -106.2646
15.7309 | -106.258
15.7347 | -106.2503
15.7384 | -106.2394
15.7422 | -106.2301
15.746 | -106.2173
15.7498 | -106.201
15.7536 | -106.1836
15.7573 | -106.1667
15.7611 | -106.1498
15.7649 | -106.137
15.7687 | -106.1255
15.7725 | -106.1128
15.7762 | -106.1004
15.78 | -106.0887
15.7838 | -106.079
15.7876 | -106.0677
15.7914 | -106.0536
15.7951 | -106.041
15.7989 | -106.0335
15.8027 | -106.0256
15.8065 | -106.018
15.8103 | -106.0182
15.814 | -106.0152
15.8178 | -106.0103
15.8216 | -106.0061
15.8254 | -106.0034
15.8292 | -106.003
15.8329 | -106.0017
15.8367 | -105.9995
15.8405 | -105.9977
15.8443 | -105.997
15.8481 | -105.9982
15.8518 | -105.9973
15.8556 | -105.9939
15.8594 | -105.9912
15.8632 | -105.9878
15.867 | -105.9838
15.8707 | -105.9804
15.8745 | -105.9773
15.8783 | -105.9735
15.8821 | -105.9698
15.8859 | -105.9648
15.8896 | -105.9584
15.8934 | -105.95
15.8972 | -105.9418
15.901 | -105.9331
15.9048 | -105.9239
15.9085 | -105.9156
15.9123 | -105.906
15.9161 | -105.8966
15.9199 | -105.8892
15.9237 | -105.8827
15.9274 | -105.8731
15.9312 | -105.863
15.935 | -105.8515
15.9388 | -105.841
15.9425 | -105.8299
15.9463 | -105.8197
15.9501 | -105.8097
15.9539 | -105.7989
15.9577 | -105.7868
15.9614 | -105.7761
15.9652 | -105.7678
15.969 | -105.759
15.9728 | -105.7501
15.9766 | -105.7439
15.9803 | -105.7393
15.9841 | -105.7339
15.9879 | -105.7309
15.9917 | -105.7276
15.9955 | -105.7237
15.9992 | -105.7195
16.003 | -105.7199
16.0068 | -105.7252
16.0106 | -105.7292
16.0144 | -105.7322
16.0181 | -105.7392
16.0219 | -105.7476
16.0257 | -105.754
16.0295 | -105.7634
16.0333 | -105.7711
16.037 | -105.7748
16.0408 | -105.7776
16.0446 | -105.7789
16.0484 | -105.7829
16.0522 | -105.7864
16.0559 | -105.7859
16.0597 | -105.786
16.0635 | -105.7848
16.0673 | -105.7807
16.0711 | -105.7732
16.0748 | -105.768
16.0786 | -105.7645
16.0824 | -105.7599
16.0862 | -105.7545
16.09 | -105.7469
16.0937 | -105.7409
16.0975 | -105.7361
16.1013 | -105.7321
16.1051 | -105.7289
16.1089 | -105.7271
16.1126 | -105.7258
16.1164 | -105.7251
16.1202 | -105.7259
16.124 | -105.729
16.1278 | -105.7333
16.1315 | -105.7408
16.1353 | -105.7496
16.1391 | -105.7589
16.1429 | -105.7698
16.1467 | -105.7816
16.1504 | -105.7937
16.1542 | -105.8078
16.158 | -105.8242
16.1618 | -105.842
16.1655 | -105.8615
16.1693 | -105.883
16.1731 | -105.9061
16.1769 | -105.9285
16.1807 | -105.9502
16.1844 | -105.9722
16.1882 | -105.9954
16.192 | -106.0151
16.1958 | -106.0336
16.1996 | -106.0549
16.2033 | -106.0773
16.2071 | -106.0993
16.2109 | -106.1222
16.2147 | -106.1446
16.2185 | -106.1647
16.2222 | -106.1839
16.226 | -106.2026
16.2298 | -106.2211
16.2336 | -106.2391
16.2374 | -106.2571
16.2411 | -106.2762
16.2449 | -106.2973
16.2487 | -106.3141
16.2525 | -106.3298
16.2563 | -106.3447
16.26 | -106.3594
16.2638 | -106.3767
16.2676 | -106.3934
16.2714 | -106.4082
16.2752 | -106.4236
16.2789 | -106.4374
16.2827 | -106.451
16.2865 | -106.4649
16.2903 | -106.4762
16.2941 | -106.4901
16.2978 | -106.5041
16.3016 | -106.5178
16.3054 | -106.526
16.3092 | -106.5316
16.313 | -106.5368
16.3167 | -106.5429
16.3205 | -106.5515
16.3243 | -106.5579
16.3281 | -106.561
16.3319 | -106.5647
16.3356 | -106.5656
16.3394 | -106.5658
16.3432 | -106.5676
16.347 | -106.567
16.3508 | -106.5644
16.3545 | -106.5619
16.3583 | -106.5526
16.3621 | -106.5497
16.3659 | -106.542
16.3696 | -106.5334
16.3734 | -106.5266
16.3772 | -106.5189
16.381 | -106.5133
16.3848 | -106.5047
16.3885 | -106.4941
16.3923 | -106.4813
16.3961 | -106.4706
16.3999 | -106.4601
16.4037 | -106.4476
16.4074 | -106.4354
16.4112 | -106.424
16.415 | -106.4144
16.4188 | -106.4055
16.4226 | -106.3972
16.4263 | -106.3927
16.4301 | -106.3881
16.4339 | -106.3838
16.4377 | -106.3771
16.4415 | -106.3709
16.4452 | -106.3675
16.449 | -106.3627
16.4528 | -106.36
16.4566 | -106.3601
16.4604 | -106.361
16.4641 | -106.3618
16.4679 | -106.3608
16.4717 | -106.357
16.4755 | -106.3554
16.4793 | -106.3575
16.483 | -106.361
16.4868 | -106.3611
16.4906 | -106.3612
16.4944 | -106.3617
16.4982 | -106.3619
16.5019 | -106.3648
16.5057 | -106.3695
16.5095 | -106.3743
16.5133 | -106.3805
16.5171 | -106.3856
16.5208 | -106.3905
16.5246 | -106.3953
16.5284 | -106.3986
16.5322 | -106.4028
16.536 | -106.4085
16.5397 | -106.4139
16.5435 | -106.4195
16.5473 | -106.4272
16.5511 | -106.4353
16.5549 | -106.4399
16.5586 | -106.4435
16.5624 | -106.449
16.5662 | -106.4532
16.57 | -106.4574
16.5738 | -106.4625
16.5775 | -106.4676
16.5813 | -106.4751
16.5851 | -106.4841
16.5889 | -106.4915
16.5926 | -106.499
16.5964 | -106.5055
16.6002 | -106.5126
16.604 | -106.5217
16.6078 | -106.5315
16.6115 | -106.5399
16.6153 | -106.549
16.6191 | -106.5557
16.6229 | -106.5662
16.6267 | -106.5764
16.6304 | -106.5888
16.6342 | -106.5973
16.638 | -106.6049
16.6418 | -106.6148
16.6456 | -106.6237
16.6493 | -106.6316
16.6531 | -106.6359
16.6569 | -106.6394
16.6607 | -106.6421
16.6645 | -106.6431
16.6682 | -106.6457
16.672 | -106.6434
16.6758 | -106.642
16.6796 | -106.639
16.6834 | -106.6323
16.6871 | -106.6268
16.6909 | -106.6184
16.6947 | -106.6063
16.6985 | -106.5931
16.7023 | -106.5843
16.706 | -106.5739
16.7098 | -106.563
16.7136 | -106.5493
16.7174 | -106.533
16.7212 | -106.5166
16.7249 | -106.5037
16.7287 | -106.4921
16.7325 | -106.4784
16.7363 | -106.4614
16.7401 | -106.4425
16.7438 | -106.4253
16.7476 | -106.405
16.7514 | -106.3843
16.7552 | -106.3645
16.759 | -106.3499
16.7627 | -106.3365
16.7665 | -106.3231
16.7703 | -106.31
16.7741 | -106.297
16.7779 | -106.2861
16.7816 | -106.2753
16.7854 | -106.2654
16.7892 | -106.256
16.793 | -106.2472
16.7967 | -106.2408
16.8005 | -106.238
16.8043 | -106.2371
16.8081 | -106.2385
16.8119 | -106.2402
16.8156 | -106.2393
16.8194 | -106.2399
16.8232 | -106.2448
16.827 | -106.2544
16.8308 | -106.2648
16.8345 | -106.277
16.8383 | -106.2909
16.8421 | -106.3041
16.8459 | -106.319
16.8497 | -106.3338
16.8534 | -106.349
16.8572 | -106.3679
16.861 | -106.3878
16.8648 | -106.4076
16.8686 | -106.4267
16.8723 | -106.4464
16.8761 | -106.466
16.8799 | -106.487
16.8837 | -106.5089
16.8875 | -106.5287
16.8912 | -106.551
16.895 | -106.5726
16.8988 | -106.5917
16.9026 | -106.6091
16.9064 | -106.6244
16.9101 | -106.6389
16.9139 | -106.6543
16.9177 | -106.6694
16.9215 | -106.6858
16.9253 | -106.6987
16.929 | -106.7116
16.9328 | -106.7229
16.9366 | -106.7304
16.9404 | -106.7363
16.9442 | -106.7397
16.9479 | -106.7442
16.9517 | -106.747
16.9555 | -106.7473
16.9593 | -106.745
16.9631 | -106.7426
16.9668 | -106.7379
16.9706 | -106.7313
16.9744 | -106.7252
16.9782 | -106.7209
16.982 | -106.7124
16.9857 | -106.7002
16.9895 | -106.6884
16.9933 | -106.6757
16.9971 | -106.6635
17.0008 | -106.6472
17.0046 | -106.6317
17.0084 | -106.6156
17.0122 | -106.5987
17.016 | -106.5837
17.0197 | -106.5687
17.0235 | -106.5519
17.0273 | -106.5349
17.0311 | -106.5186
17.0349 | -106.5045
17.0386 | -106.4898
17.0424 | -106.4724
17.0462 | -106.4548
17.05 | -106.44
17.0538 | -106.4271
17.0575 | -106.4168
17.0613 | -106.4048
17.0651 | -106.3895
17.0689 | -106.3761
17.0727 | -106.3633
17.0764 | -106.3518
17.0802 | -106.3421
17.084 | -106.3369
17.0878 | -106.3302
17.0916 | -106.324
17.0953 | -106.3217
17.0991 | -106.3195
17.1029 | -106.3158
17.1067 | -106.3138
17.1105 | -106.314
17.1142 | -106.3152
17.118 | -106.3185
17.1218 | -106.3201
17.1256 | -106.3243
17.1294 | -106.3315
17.1331 | -106.3391
17.1369 | -106.349
17.1407 | -106.3601
17.1445 | -106.3695
17.1483 | -106.3795
17.152 | -106.3929
17.1558 | -106.4057
17.1596 | -106.4187
17.1634 | -106.4316
17.1672 | -106.4428
17.1709 | -106.4558
17.1747 | -106.4678
17.1785 | -106.4801
17.1823 | -106.4928
17.1861 | -106.5051
17.1898 | -106.5159
17.1936 | -106.5262
17.1974 | -106.5372
17.2012 | -106.5489
17.205 | -106.5607
17.2087 | -106.572
17.2125 | -106.5816
17.2163 | -106.5891
17.2201 | -106.596
17.2238 | -106.6038
17.2276 | -106.6142
17.2314 | -106.6222
17.2352 | -106.6292
17.239 | -106.6371
17.2427 | -106.6423
17.2465 | -106.647
17.2503 | -106.6493
17.2541 | -106.6551
17.2579 | -106.6619
17.2616 | -106.6661
17.2654 | -106.6703
17.2692 | -106.6774
17.273 | -106.6854
17.2768 | -106.6912
17.2805 | -106.6943
17.2843 | -106.6973
17.2881 | -106.7
17.2919 | -106.7058
17.2957 | -106.709
17.2994 | -106.71
17.3032 | -106.7114
17.307 | -106.7142
17.3108 | -106.7195
17.3146 | -106.721
17.3183 | -106.7228
17.3221 | -106.7248
17.3259 | -106.7232
17.3297 | -106.7201
17.3335 | -106.7181
17.3372 | -106.7153
17.341 | -106.7065
17.3448 | -106.6967
17.3486 | -106.687
17.3524 | -106.6777
17.3561 | -106.669
17.3599 | -106.6581
17.3637 | -106.6462
17.3675 | -106.6352
17.3713 | -106.621
17.375 | -106.6055
17.3788 | -106.5907
17.3826 | -106.5719
17.3864 | -106.5525
17.3902 | -106.536
17.3939 | -106.5206
17.3977 | -106.5016
17.4015 | -106.4822
17.4053 | -106.4617
17.4091 | -106.4417
17.4128 | -106.4219
17.4166 | -106.4031
17.4204 | -106.3848
17.4242 | -106.365
17.4279 | -106.3487
17.4317 | -106.3331
17.4355 | -106.3153
17.4393 | -106.2991
17.4431 | -106.2841
17.4468 | -106.279
17.4506 | -106.2678
17.4544 | -106.2586
17.4582 | -106.2495
17.462 | -106.2424
17.4657 | -106.2378
17.4695 | -106.2345
17.4733 | -106.2317
17.4771 | -106.2323
17.4809 | -106.2356
17.4846 | -106.2394
17.4884 | -106.243
17.4922 | -106.2471
17.496 | -106.2539
17.4998 | -106.2613
17.5035 | -106.269
17.5073 | -106.2766
17.5111 | -106.2853
17.5149 | -106.2952
17.5187 | -106.3053
17.5224 | -106.3149
17.5262 | -106.3288
17.53 | -106.3453
17.5338 | -106.3618
17.5376 | -106.3761
17.5413 | -106.3872
17.5451 | -106.4011
17.5489 | -106.4159
17.5527 | -106.4306
17.5565 | -106.4452
17.5602 | -106.4577
17.564 | -106.4715
17.5678 | -106.4858
17.5716 | -106.4989
17.5754 | -106.513
17.5791 | -106.5236
17.5829 | -106.5361
17.5867 | -106.5484
17.5905 | -106.5624
17.5943 | -106.5753
17.598 | -106.5882
17.6018 | -106.5982
17.6056 | -106.6048
17.6094 | -106.6132
17.6132 | -106.6206
17.6169 | -106.6276
17.6207 | -106.6335
17.6245 | -106.6401
17.6283 | -106.6458
17.6321 | -106.6521
17.6358 | -106.6578
17.6396 | -106.6614
17.6434 | -106.6624
17.6472 | -106.6654
17.651 | -106.668
17.6547 | -106.6713
17.6585 | -106.6714
17.6623 | -106.6713
17.6661 | -106.6713
17.6698 | -106.6697
17.6736 | -106.6704
17.6774 | -106.6681
17.6812 | -106.6663
17.685 | -106.6664
17.6887 | -106.6651
17.6925 | -106.6644
17.6963 | -106.6628
17.7001 | -106.6609
17.7039 | -106.6595
17.7076 | -106.6572
17.7114 | -106.6553
17.7152 | -106.6508
17.719 | -106.6446
17.7228 | -106.6413
17.7265 | -106.6354
17.7303 | -106.6264
17.7341 | -106.6165
17.7379 | -106.6088
17.7417 | -106.6027
17.7454 | -106.5978
17.7492 | -106.5884
17.753 | -106.5801
17.7568 | -106.5684
17.7606 | -106.5548
17.7643 | -106.5442
17.7681 | -106.5328
17.7719 | -106.5193
17.7757 | -106.5085
17.7795 | -106.4963
17.7832 | -106.4833
17.787 | -106.4723
17.7908 | -106.4626
17.7946 | -106.4517
17.7984 | -106.4393
17.8021 | -106.4289
17.8059 | -106.4177
17.8097 | -106.4042
17.8135 | -106.3922
17.8173 | -106.3833
17.821 | -106.3751
17.8248 | -106.3666
17.8286 | -106.3568
17.8324 | -106.3489
17.8362 | -106.3399
17.8399 | -106.33
17.8437 | -106.3216
17.8475 | -106.313
17.8513 | -106.303
17.8551 | -106.2922
17.8588 | -106.2819
17.8626 | -106.2716
17.8664 | -106.2634
17.8702 | -106.2565
17.8739 | -106.2507
17.8777 | -106.2437
17.8815 | -106.2344
17.8853 | -106.2271
17.8891 | -106.2227
17.8928 | -106.2185
17.8966 | -106.2161
17.9004 | -106.2147
17.9042 | -106.213
17.908 | -106.2091
17.9117 | -106.207
17.9155 | -106.2056
17.9193 | -106.2031
17.9231 | -106.2007
17.9269 | -106.199
17.9306 | -106.1987
17.9344 | -106.1978
17.9382 | -106.196
17.942 | -106.1949
17.9458 | -106.1919
17.9495 | -106.1893
17.9533 | -106.189
17.9571 | -106.1888
17.9609 | -106.1894
17.9647 | -106.1868
17.9684 | -106.1861
17.9722 | -106.1864
17.976 | -106.1865
17.9798 | -106.1885
17.9836 | -106.1919
17.9873 | -106.1957
17.9911 | -106.1977
17.9949 | -106.1997
17.9987 | -106.2012
18.0025 | -106.2035
18.0062 | -106.2083
18.01 | -106.2102
18.0138 | -106.2143
18.0176 | -106.222
18.0214 | -106.2239
18.0251 | -106.2249
18.0289 | -106.2292
18.0327 | -106.2324
18.0365 | -106.2353
18.0403 | -106.2372
18.044 | -106.2378
18.0478 | -106.2392
18.0516 | -106.2417
18.0554 | -106.2433
18.0592 | -106.2435
18.0629 | -106.2447
18.0667 | -106.2453
18.0705 | -106.2442
18.0743 | -106.24
18.0781 | -106.2363
18.0818 | -106.2333
18.0856 | -106.2309
18.0894 | -106.2304
18.0932 | -106.2281
18.0969 | -106.2266
18.1007 | -106.225
18.1045 | -106.2244
18.1083 | -106.2245
18.1121 | -106.2258
18.1158 | -106.2281
18.1196 | -106.2307
18.1234 | -106.2315
18.1272 | -106.2326
18.131 | -106.2355
18.1347 | -106.2396
18.1385 | -106.2468
18.1423 | -106.2546
18.1461 | -106.2628
18.1499 | -106.2714
18.1536 | -106.2813
18.1574 | -106.2937
18.1612 | -106.307
18.165 | -106.3229
18.1688 | -106.339
18.1725 | -106.3549
18.1763 | -106.3737
18.1801 | -106.3914
18.1839 | -106.4065
18.1877 | -106.4233
18.1914 | -106.4412
18.1952 | -106.4588
18.199 | -106.4777
18.2028 | -106.4978
18.2066 | -106.5169
18.2103 | -106.5332
18.2141 | -106.5445
18.2179 | -106.5545
18.2217 | -106.5684
18.2255 | -106.5825
18.2292 | -106.5933
18.233 | -106.5981
18.2368 | -106.6029
18.2406 | -106.6096
18.2444 | -106.6157
18.2481 | -106.6208
18.2519 | -106.6237
18.2557 | -106.6241
18.2595 | -106.6239
18.2633 | -106.6229
18.267 | -106.6177
18.2708 | -106.616
18.2746 | -106.6142
18.2784 | -106.6123
18.2822 | -106.612
18.2859 | -106.6097
18.2897 | -106.6053
18.2935 | -106.601
18.2973 | -106.5963
18.301 | -106.5916
18.3048 | -106.5872
18.3086 | -106.5836
18.3124 | -106.5805
18.3162 | -106.5754
18.3199 | -106.5696
18.3237 | -106.5638
18.3275 | -106.5608
18.3313 | -106.5574
18.3351 | -106.5551
18.3388 | -106.5504
18.3426 | -106.5446
18.3464 | -106.5397
18.3502 | -106.5322
18.354 | -106.5269
18.3577 | -106.5208
18.3615 | -106.5117
18.3653 | -106.5045
18.3691 | -106.4983
18.3729 | -106.4911
18.3766 | -106.4837
18.3804 | -106.4746
18.3842 | -106.4669
18.388 | -106.4563
18.3918 | -106.4443
18.3955 | -106.4343
18.3993 | -106.4257
18.4031 | -106.4164
18.4069 | -106.4025
18.4107 | -106.3878
18.4144 | -106.3751
18.4182 | -106.3613
18.422 | -106.347
18.4258 | -106.3353
18.4296 | -106.3251
18.4333 | -106.3136
18.4371 | -106.2999
18.4409 | -106.2897
18.4447 | -106.2809
18.4485 | -106.2715
18.4522 | -106.2612
18.456 | -106.2541
18.4598 | -106.2464
18.4636 | -106.2396
18.4674 | -106.2343
18.4711 | -106.2326
18.4749 | -106.2343
18.4787 | -106.2374
18.4825 | -106.2406
18.4863 | -106.2409
18.49 | -106.2434
18.4938 | -106.2505
18.4976 | -106.2571
18.5014 | -106.2632
18.5052 | -106.2697
18.5089 | -106.2795
18.5127 | -106.2925
18.5165 | -106.3069
18.5203 | -106.3195
18.524 | -106.3351
18.5278 | -106.3542
18.5316 | -106.3722
18.5354 | -106.3892
18.5392 | -106.4081
18.5429 | -106.4233
18.5467 | -106.4392
18.5505 | -106.4574
18.5543 | -106.4737
18.5581 | -106.49
18.5618 | -106.5079
18.5656 | -106.5246
18.5694 | -106.5391
18.5732 | -106.5517
18.577 | -106.5641
18.5807 | -106.578
18.5845 | -106.5893
18.5883 | -106.6006
18.5921 | -106.6128
18.5959 | -106.6239
18.5996 | -106.6331
18.6034 | -106.6401
18.6072 | -106.6488
18.611 | -106.6571
18.6148 | -106.6646
18.6185 | -106.6697
18.6223 | -106.6739
18.6261 | -106.6805
18.6299 | -106.6844
18.6337 | -106.6873
18.6374 | -106.69
18.6412 | -106.6923
18.645 | -106.696
18.6488 | -106.6985
18.6526 | -106.7033
18.6563 | -106.7076
18.6601 | -106.7094
18.6639 | -106.7087
18.6677 | -106.7083
18.6715 | -106.7105
18.6752 | -106.7111
18.679 | -106.7101
18.6828 | -106.7102
18.6866 | -106.7087
18.6904 | -106.7039
18.6941 | -106.6998
18.6979 | -106.6978
18.7017 | -106.6966
18.7055 | -106.6955
18.7093 | -106.6926
18.713 | -106.6886
18.7168 | -106.6854
18.7206 | -106.6789
18.7244 | -106.6727
18.7281 | -106.6689
18.7319 | -106.6645
18.7357 | -106.6591
18.7395 | -106.6547
18.7433 | -106.65
18.747 | -106.6452
18.7508 | -106.6439
18.7546 | -106.6435
18.7584 | -106.6413
18.7622 | -106.6406
18.7659 | -106.6383
18.7697 | -106.6368
18.7735 | -106.6366
18.7773 | -106.6359
18.7811 | -106.636
18.7848 | -106.6379
18.7886 | -106.6389
18.7924 | -106.6403
18.7962 | -106.6408
18.8 | -106.6398
18.8037 | -106.6387
18.8075 | -106.6398
18.8113 | -106.6428
18.8151 | -106.6462
18.8189 | -106.6475
18.8226 | -106.6493
18.8264 | -106.6489
18.8302 | -106.6489
18.834 | -106.6514
18.8378 | -106.6532
18.8415 | -106.6541
18.8453 | -106.6534
18.8491 | -106.6518
18.8529 | -106.6508
18.8567 | -106.6478
18.8604 | -106.6397
18.8642 | -106.6333
18.868 | -106.628
18.8718 | -106.6236
18.8756 | -106.6176
18.8793 | -106.6106
18.8831 | -106.6039
18.8869 | -106.5981
18.8907 | -106.5908
18.8945 | -106.5838
18.8982 | -106.5777
18.902 | -106.5704
18.9058 | -106.5626
18.9096 | -106.555
18.9134 | -106.5479
18.9171 | -106.5408
18.9209 | -106.5336
18.9247 | -106.5276
18.9285 | -106.5213
18.9322 | -106.5164
18.936 | -106.5123
18.9398 | -106.5096
18.9436 | -106.5073
18.9474 | -106.5038
18.9511 | -106.4993
18.9549 | -106.4982
18.9587 | -106.4974
18.9625 | -106.4936
18.9663 | -106.4912
18.97 | -106.4881
18.9738 | -106.4871
18.9776 | -106.4868
18.9814 | -106.4865
18.9852 | -106.4885
18.9889 | -106.4935
18.9927 | -106.4944
18.9965 | -106.492
19.0003 | -106.4912
19.0041 | -106.4907
19.0078 | -106.4909
19.0116 | -106.4909
19.0154 | -106.49
19.0192 | -106.4904
19.023 | -106.4861
19.0267 | -106.4812
19.0305 | -106.4755
19.0343 | -106.4698
19.0381 | -106.4656
19.0419 | -106.4565
19.0456 | -106.4476
19.0494 | -106.4398
19.0532 | -106.4317
19.057 | -106.4236
19.0608 | -106.4163
19.0645 | -106.4083
19.0683 | -106.4003
19.0721 | -106.3928
19.0759 | -106.383
19.0797 | -106.3727
19.0834 | -106.3677
19.0872 | -106.3587
19.091 | -106.3528
19.0948 | -106.3481
19.0986 | -106.3441
19.1023 | -106.3375
19.1061 | -106.3326
19.1099 | -106.3296
19.1137 | -106.3267
19.1175 | -106.3236
19.1212 | -106.322
19.125 | -106.3217
19.1288 | -106.3229
19.1326 | -106.3255
19.1364 | -106.3258
19.1401 | -106.3239
19.1439 | -106.3255
19.1477 | -106.3283
19.1515 | -106.331
19.1552 | -106.3344
19.159 | -106.3366
19.1628 | -106.3413
19.1666 | -106.3436
19.1704 | -106.3456
19.1741 | -106.3473
19.1779 | -106.3479
19.1817 | -106.3504
19.1855 | -106.3513
19.1893 | -106.3512
19.193 | -106.3503
19.1968 | -106.352
19.2006 | -106.3521
19.2044 | -106.3501
19.2082 | -106.347
19.2119 | -106.3413
19.2157 | -106.3348
19.2195 | -106.3276
19.2233 | -106.321
19.2271 | -106.3123
19.2308 | -106.3043
19.2346 | -106.2967
19.2384 | -106.2871
19.2422 | -106.2781
19.246 | -106.2656
19.2497 | -106.255
19.2535 | -106.2432
19.2573 | -106.2298
19.2611 | -106.218
19.2649 | -106.2065
19.2686 | -106.1971
19.2724 | -106.1869
19.2762 | -106.1767
19.28 | -106.1665
19.2838 | -106.1569
19.2875 | -106.1477
19.2913 | -106.1382
19.2951 | -106.1344
19.2989 | -106.1317
19.3027 | -106.1304
19.3064 | -106.1289
19.3102 | -106.1271
19.314 | -106.1288
19.3178 | -106.1291
19.3216 | -106.1274
19.3253 | -106.1282
19.3291 | -106.1286
19.3329 | -106.1305
19.3367 | -106.1334
19.3405 | -106.1383
19.3442 | -106.1432
19.348 | -106.1475
19.3518 | -106.1532
19.3556 | -106.1597
19.3593 | -106.1654
19.3631 | -106.1709
19.3669 | -106.1754
19.3707 | -106.1814
19.3745 | -106.1864
19.3782 | -106.1903
19.382 | -106.1939
19.3858 | -106.1993
19.3896 | -106.2056
19.3934 | -106.2117
19.3971 | -106.2166
19.4009 | -106.2201
19.4047 | -106.2235
19.4085 | -106.2294
19.4123 | -106.2341
19.416 | -106.2396
19.4198 | -106.2441
19.4236 | -106.2466
19.4274 | -106.2489
19.4312 | -106.2496
19.4349 | -106.2511
19.4387 | -106.2535
19.4425 | -106.2574
19.4463 | -106.2598
19.4501 | -106.2642
19.4538 | -106.2692
19.4576 | -106.2724
19.4614 | -106.2772
19.4652 | -106.2828
19.469 | -106.2883
19.4727 | -106.291
19.4765 | -106.2952
19.4803 | -106.3014
19.4841 | -106.3077
19.4879 | -106.3152
19.4916 | -106.3239
19.4954 | -106.3323
19.4992 | -106.3378
19.503 | -106.3427
19.5068 | -106.3469
19.5105 | -106.3494
19.5143 | -106.3541
19.5181 | -106.3597
19.5219 | -106.3643
19.5257 | -106.3682
19.5294 | -106.3729
19.5332 | -106.3752
19.537 | -106.3761
19.5408 | -106.3799
19.5446 | -106.3834
19.5483 | -106.3877
19.5521 | -106.3917
19.5559 | -106.3942
19.5597 | -106.3987
19.5635 | -106.4027
19.5672 | -106.406
19.571 | -106.4066
19.5748 | -106.4073
19.5786 | -106.4077
19.5823 | -106.4098
19.5861 | -106.4104
19.5899 | -106.4112
19.5937 | -106.4119
19.5975 | -106.4149
19.6012 | -106.4159
19.605 | -106.4168
19.6088 | -106.4164
19.6126 | -106.4162
19.6164 | -106.4152
19.6201 | -106.4138
19.6239 | -106.4133
19.6277 | -106.4112
19.6315 | -106.4086
19.6353 | -106.407
19.639 | -106.4052
19.6428 | -106.4023
19.6466 | -106.3973
19.6504 | -106.3904
19.6542 | -106.3867
19.6579 | -106.3859
19.6617 | -106.3842
19.6655 | -106.3801
19.6693 | -106.3763
19.6731 | -106.3718
19.6768 | -106.3695
19.6806 | -106.3674
19.6844 | -106.3641
19.6882 | -106.3607
19.692 | -106.3592
19.6957 | -106.3569
19.6995 | -106.3547
19.7033 | -106.3539
19.7071 | -106.3505
19.7109 | -106.3493
19.7146 | -106.3494
19.7184 | -106.3464
19.7222 | -106.343
19.726 | -106.3403
19.7298 | -106.3373
19.7335 | -106.3361
19.7373 | -106.3349
19.7411 | -106.3333
19.7449 | -106.3323
19.7487 | -106.3328
19.7524 | -106.3335
19.7562 | -106.3349
19.76 | -106.3374
19.7638 | -106.3396
19.7676 | -106.3425
19.7713 | -106.3458
19.7751 | -106.3497
19.7789 | -106.3547
19.7827 | -106.3605
19.7864 | -106.3668
19.7902 | -106.3749
19.794 | -106.3801
19.7978 | -106.3863
19.8016 | -106.3945
19.8053 | -106.4017
19.8091 | -106.4094
19.8129 | -106.4183
19.8167 | -106.4285
19.8205 | -106.438
19.8242 | -106.4456
19.828 | -106.4523
19.8318 | -106.4597
19.8356 | -106.4686
19.8394 | -106.4755
19.8431 | -106.4834
19.8469 | -106.4913
19.8507 | -106.5007
19.8545 | -106.5104
19.8583 | -106.5177
19.862 | -106.525
19.8658 | -106.534
19.8696 | -106.5416
19.8734 | -106.5501
19.8772 | -106.5564
19.8809 | -106.5619
19.8847 | -106.5679
19.8885 | -106.5725
19.8923 | -106.5773
19.8961 | -106.5825
19.8998 | -106.5857
19.9036 | -106.5897
19.9074 | -106.5922
19.9112 | -106.5951
19.915 | -106.5965
19.9187 | -106.5962
19.9225 | -106.5976
19.9263 | -106.5994
19.9301 | -106.6011
19.9339 | -106.6028
19.9376 | -106.6009
19.9414 | -106.599
19.9452 | -106.5959
19.949 | -106.5917
19.9528 | -106.5864
19.9565 | -106.5835
19.9603 | -106.5782
19.9641 | -106.5725
19.9679 | -106.5678
19.9717 | -106.5637
19.9754 | -106.561
19.9792 | -106.5569
19.983 | -106.5515
19.9868 | -106.5453
19.9906 | -106.5404
19.9943 | -106.5362
19.9981 | -106.5313
20.0019 | -106.5275
20.0057 | -106.5235
20.0094 | -106.5193
20.0132 | -106.5137
20.017 | -106.5073
20.0208 | -106.5006
20.0246 | -106.4955
20.0283 | -106.4894
20.0321 | -106.4838
20.0359 | -106.4797
20.0397 | -106.4756
20.0435 | -106.4716
20.0472 | -106.4684
20.051 | -106.4651
20.0548 | -106.462
20.0586 | -106.4585
20.0624 | -106.458
20.0661 | -106.4577
20.0699 | -106.4574
20.0737 | -106.4563
20.0775 | -106.4548
20.0813 | -106.4507
20.085 | -106.4484
20.0888 | -106.4476
20.0926 | -106.4463
20.0964 | -106.4469
20.1002 | -106.4441
20.1039 | -106.4429
20.1077 | -106.4436
20.1115 | -106.4444
20.1153 | -106.4447
20.1191 | -106.4456
20.1228 | -106.4449
20.1266 | -106.4436
20.1304 | -106.4435
20.1342 | -106.4428
20.138 | -106.4418
20.1417 | -106.4419
20.1455 | -106.4407
20.1493 | -106.4385
20.1531 | -106.4366
20.1569 | -106.4344
20.1606 | -106.4337
20.1644 | -106.4321
20.1682 | -106.4304
20.172 | -106.4279
20.1758 | -106.4266
20.1795 | -106.4255
20.1833 | -106.4228
20.1871 | -106.4216
20.1909 | -106.4174
20.1947 | -106.4119
20.1984 | -106.4103
20.2022 | -106.4103
20.206 | -106.4101
20.2098 | -106.4089
20.2136 | -106.4071
20.2173 | -106.4078
20.2211 | -106.4073
20.2249 | -106.4087
20.2287 | -106.409
20.2324 | -106.4065
20.2362 | -106.4071
20.24 | -106.4101
20.2438 | -106.4124
20.2476 | -106.4148
20.2513 | -106.4198
20.2551 | -106.4239
20.2589 | -106.4252
20.2627 | -106.427
20.2665 | -106.4292
20.2702 | -106.4295
20.274 | -106.4333
20.2778 | -106.4407
20.2816 | -106.4458
20.2854 | -106.449
20.2891 | -106.4539
20.2929 | -106.4587
20.2967 | -106.4639
20.3005 | -106.4697
20.3043 | -106.4748
20.308 | -106.4787
20.3118 | -106.4823
20.3156 | -106.4866
20.3194 | -106.4918
20.3232 | -106.4972
20.3269 | -106.5006
20.3307 | -106.5023
20.3345 | -106.5035
20.3383 | -106.5049
20.3421 | -106.5062
20.3458 | -106.5089
20.3496 | -106.5103
20.3534 | -106.5111
20.3572 | -106.5115
20.361 | -106.5127
20.3647 | -106.515
20.3685 | -106.5194
20.3723 | -106.5229
20.3761 | -106.5233
20.3799 | -106.5223
20.3836 | -106.5183
20.3874 | -106.5168
20.3912 | -106.5143
20.395 | -106.5115
20.3988 | -106.5108
20.4025 | -106.509
20.4063 | -106.5071
20.4101 | -106.5066
20.4139 | -106.506
20.4177 | -106.5011
20.4214 | -106.4955
20.4252 | -106.4928
20.429 | -106.4885
20.4328 | -106.4834
20.4365 | -106.4794
20.4403 | -106.4772
20.4441 | -106.4744
20.4479 | -106.4696
20.4517 | -106.4666
20.4554 | -106.4636
20.4592 | -106.4601
20.463 | -106.4561
20.4668 | -106.4541
20.4706 | -106.4509
20.4743 | -106.4479
20.4781 | -106.4449
20.4819 | -106.4429
20.4857 | -106.4427
20.4895 | -106.4387
20.4932 | -106.4346
20.497 | -106.4305
20.5008 | -106.4282
20.5046 | -106.4265
20.5084 | -106.4213
20.5121 | -106.4186
20.5159 | -106.4141
20.5197 | -106.4104
20.5235 | -106.4106
20.5273 | -106.4083
20.531 | -106.4048
20.5348 | -106.4004
20.5386 | -106.3968
20.5424 | -106.3922
20.5462 | -106.3881
20.5499 | -106.3851
20.5537 | -106.3815
20.5575 | -106.3785
20.5613 | -106.377
20.5651 | -106.3717
20.5688 | -106.3651
20.5726 | -106.3604
20.5764 | -106.3574
20.5802 | -106.3521
20.584 | -106.3463
20.5877 | -106.3422
20.5915 | -106.3386
20.5953 | -106.335
20.5991 | -106.3307
20.6029 | -106.326
20.6066 | -106.321
20.6104 | -106.3159
20.6142 | -106.3096
20.618 | -106.3047
20.6218 | -106.3013
20.6255 | -106.2984
20.6293 | -106.2938
20.6331 | -106.29
20.6369 | -106.287
20.6407 | -106.2821
20.6444 | -106.2766
20.6482 | -106.2729
20.652 | -106.27
20.6558 | -106.2656
20.6595 | -106.2593
20.6633 | -106.253
20.6671 | -106.249
20.6709 | -106.2462
20.6747 | -106.2428
20.6784 | -106.2362
20.6822 | -106.2299
20.686 | -106.2218
20.6898 | -106.2131
20.6936 | -106.2079
20.6973 | -106.2017
20.7011 | -106.1958
20.7049 | -106.189
20.7087 | -106.1848
20.7125 | -106.1806
20.7162 | -106.1722
20.72 | -106.1666
20.7238 | -106.1621
20.7276 | -106.1576
20.7314 | -106.153
20.7351 | -106.1486
20.7389 | -106.1446
20.7427 | -106.1404
20.7465 | -106.1377
20.7503 | -106.1354
20.754 | -106.1334
20.7578 | -106.1319
20.7616 | -106.1285
20.7654 | -106.1257
20.7692 | -106.1231
20.7729 | -106.121
20.7767 | -106.1205
20.7805 | -106.1187
20.7843 | -106.1176
20.7881 | -106.1156
20.7918 | -106.1131
20.7956 | -106.1109
20.7994 | -106.109
20.8032 | -106.1064
20.807 | -106.1035
20.8107 | -106.1017
20.8145 | -106.1
20.8183 | -106.0978
20.8221 | -106.0964
20.8259 | -106.0944
20.8296 | -106.0936
20.8334 | -106.0925
20.8372 | -106.0908
20.841 | -106.0877
20.8448 | -106.0856
20.8485 | -106.0839
20.8523 | -106.0815
20.8561 | -106.0798
20.8599 | -106.0783
20.8636 | -106.0755
20.8674 | -106.0711
20.8712 | -106.069
20.875 | -106.0678
20.8788 | -106.0679
20.8825 | -106.0684
20.8863 | -106.067
20.8901 | -106.0666
20.8939 | -106.0669
20.8977 | -106.0675
20.9014 | -106.0698
20.9052 | -106.0712
20.909 | -106.0731
20.9128 | -106.0763
20.9166 | -106.079
20.9203 | -106.0832
20.9241 | -106.0883
20.9279 | -106.0921
20.9317 | -106.0948
20.9355 | -106.1012
20.9392 | -106.1091
20.943 | -106.1144
20.9468 | -106.1208
20.9506 | -106.1282
20.9544 | -106.1376
20.9581 | -106.1479
20.9619 | -106.1563
20.9657 | -106.1666
20.9695 | -106.178
20.9733 | -106.1907
20.977 | -106.2033
20.9808 | -106.2165
20.9846 | -106.2286
20.9884 | -106.2394
20.9922 | -106.2504
20.9959 | -106.2626
20.9997 | -106.2741
21.0035 | -106.2862
21.0073 | -106.2982
21.0111 | -106.3121
21.0148 | -106.3253
21.0186 | -106.3355
21.0224 | -106.3444
21.0262 | -106.3553
21.03 | -106.3648
21.0337 | -106.3726
21.0375 | -106.3818
21.0413 | -106.392
21.0451 | -106.3997
21.0489 | -106.4054
21.0526 | -106.4108
21.0564 | -106.4175
21.0602 | -106.4251
21.064 | -106.4313
21.0678 | -106.4371
21.0715 | -106.4421
21.0753 | -106.4453
21.0791 | -106.4479
21.0829 | -106.4506
21.0866 | -106.4518
21.0904 | -106.4533
21.0942 | -106.4522
21.098 | -106.4512
21.1018 | -106.4536
21.1055 | -106.4521
21.1093 | -106.449
21.1131 | -106.4465
21.1169 | -106.4434
21.1207 | -106.4415
21.1244 | -106.4411
21.1282 | -106.4383
21.132 | -106.4329
21.1358 | -106.4287
21.1396 | -106.4279
21.1433 | -106.4274
21.1471 | -106.4255
21.1509 | -106.423
21.1547 | -106.421
21.1585 | -106.4185
21.1622 | -106.4162
21.166 | -106.4153
21.1698 | -106.4118
21.1736 | -106.4088
21.1774 | -106.4057
21.1811 | -106.4014
21.1849 | -106.3994
21.1887 | -106.398
21.1925 | -106.3964
21.1963 | -106.3937
21.2 | -106.3896
21.2038 | -106.3856
21.2076 | -106.3826
21.2114 | -106.3806
21.2152 | -106.3805
21.2189 | -106.3797
21.2227 | -106.3798
21.2265 | -106.3791
21.2303 | -106.3783
21.2341 | -106.3765
21.2378 | -106.3745
21.2416 | -106.3754
21.2454 | -106.3745
21.2492 | -106.3729
21.253 | -106.3728
21.2567 | -106.3707
21.2605 | -106.3718
21.2643 | -106.3718
21.2681 | -106.3711
21.2719 | -106.369
21.2756 | -106.368
21.2794 | -106.3663
21.2832 | -106.3655
21.287 | -106.3663
21.2907 | -106.3658
21.2945 | -106.3655
21.2983 | -106.3648
21.3021 | -106.3646
21.3059 | -106.3631
21.3096 | -106.3617
21.3134 | -106.3608
21.3172 | -106.3593
21.321 | -106.3591
21.3248 | -106.3605
21.3285 | -106.3609
21.3323 | -106.3608
21.3361 | -106.3592
21.3399 | -106.3577
21.3437 | -106.3567
21.3474 | -106.3571
21.3512 | -106.3586
21.355 | -106.3598
21.3588 | -106.3586
21.3626 | -106.3562
21.3663 | -106.3561
21.3701 | -106.3573
21.3739 | -106.3574
21.3777 | -106.3589
21.3815 | -106.3622
21.3852 | -106.3649
21.389 | -106.3683
21.3928 | -106.3707
21.3966 | -106.3735
21.4004 | -106.3761
21.4041 | -106.3765
21.4079 | -106.3763
21.4117 | -106.3771
21.4155 | -106.3767
21.4193 | -106.378
21.423 | -106.378
21.4268 | -106.3774
21.4306 | -106.3766
21.4344 | -106.3795
21.4382 | -106.3802
21.4419 | -106.3799
21.4457 | -106.3783
21.4495 | -106.3751
21.4533 | -106.373
21.4571 | -106.3701
21.4608 | -106.3684
21.4646 | -106.3662
21.4684 | -106.3643
21.4722 | -106.3622
21.476 | -106.3584
21.4797 | -106.355
21.4835 | -106.352
21.4873 | -106.3473
21.4911 | -106.3398
21.4948 | -106.3339
21.4986 | -106.328
21.5024 | -106.3238
21.5062 | -106.3205
21.51 | -106.3152
21.5137 | -106.3094
21.5175 | -106.3033
21.5213 | -106.2971
21.5251 | -106.2899
21.5289 | -106.2839
21.5326 | -106.2796
21.5364 | -106.2744
21.5402 | -106.2697
21.544 | -106.2642
21.5478 | -106.2597
21.5515 | -106.2573
21.5553 | -106.2542
21.5591 | -106.2499
21.5629 | -106.244
21.5667 | -106.2414
21.5704 | -106.2395
21.5742 | -106.2371
21.578 | -106.237
21.5818 | -106.2364
21.5856 | -106.2367
21.5893 | -106.2334
21.5931 | -106.2301
21.5969 | -106.2287
21.6007 | -106.2265
21.6045 | -106.2251
21.6082 | -106.2261
21.612 | -106.2284
21.6158 | -106.2311
21.6196 | -106.2333
21.6234 | -106.2372
21.6271 | -106.2407
21.6309 | -106.2435
21.6347 | -106.2461
21.6385 | -106.2498
21.6423 | -106.2541
21.646 | -106.2577
21.6498 | -106.2617
21.6536 | -106.2625
21.6574 | -106.2646
21.6612 | -106.268
21.6649 | -106.2718
21.6687 | -106.2766
21.6725 | -106.2792
21.6763 | -106.279
21.6801 | -106.2782
21.6838 | -106.2766
21.6876 | -106.2766
21.6914 | -106.2762
21.6952 | -106.2746
21.699 | -106.275
21.7027 | -106.2738
21.7065 | -106.2725
21.7103 | -106.2707
21.7141 | -106.269
21.7178 | -106.2646
21.7216 | -106.2596
21.7254 | -106.2536
21.7292 | -106.2486
21.733 | -106.2432
21.7367 | -106.2394
21.7405 | -106.2347
21.7443 | -106.2302
21.7481 | -106.223
21.7519 | -106.2159
21.7556 | -106.2109
21.7594 | -106.2054
21.7632 | -106.1989
21.767 | -106.1905
21.7708 | -106.1833
21.7745 | -106.1755
21.7783 | -106.1685
21.7821 | -106.164
21.7859 | -106.1589
21.7897 | -106.1526
21.7934 | -106.1452
21.7972 | -106.1403
21.801 | -106.1346
21.8048 | -106.131
21.8086 | -106.1279
21.8123 | -106.1218
21.8161 | -106.1166
21.8199 | -106.1121
21.8237 | -106.1054
21.8275 | -106.0984
21.8312 | -106.0944
21.835 | -106.0907
21.8388 | -106.0858
21.8426 | -106.0819
21.8464 | -106.0789
21.8501 | -106.0746
21.8539 | -106.0721
21.8577 | -106.07
21.8615 | -106.0666
21.8653 | -106.0619
21.869 | -106.0606
21.8728 | -106.0602
21.8766 | -106.0583
21.8804 | -106.0574
21.8842 | -106.0562
21.8879 | -106.0556
21.8917 | -106.0554
21.8955 | -106.0558
21.8993 | -106.0594
21.9031 | -106.0618
21.9068 | -106.0626
21.9106 | -106.0654
21.9144 | -106.0672
21.9182 | -106.068
21.9219 | -106.0665
21.9257 | -106.0642
21.9295 | -106.0617
21.9333 | -106.0596
21.9371 | -106.0583
21.9408 | -106.0582
21.9446 | -106.0582
21.9484 | -106.057
21.9522 | -106.0581
21.956 | -106.0596
21.9597 | -106.0591
21.9635 | -106.0607
21.9673 | -106.0606
21.9711 | -106.0574
21.9749 | -106.0542
21.9786 | -106.052
21.9824 | -106.0502
21.9862 | -106.0501
21.99 | -106.0479
21.9938 | -106.0439
21.9975 | -106.0392
22.0013 | -106.0341
22.0051 | -106.0296
22.0089 | -106.0231
22.0127 | -106.019
22.0164 | -106.0138
22.0202 | -106.0082
22.024 | -106.0019
22.0278 | -105.9966
22.0316 | -105.9907
22.0353 | -105.9838
22.0391 | -105.9773
22.0429 | -105.9708
22.0467 | -105.9636
22.0505 | -105.9558
22.0542 | -105.9479
22.058 | -105.9413
22.0618 | -105.9343
22.0656 | -105.9277
22.0694 | -105.9216
22.0731 | -105.9155
22.0769 | -105.9093
22.0807 | -105.9024
22.0845 | -105.8966
22.0883 | -105.8914
22.092 | -105.8853
22.0958 | -105.8801
22.0996 | -105.8764
22.1034 | -105.8731
22.1072 | -105.8699
22.1109 | -105.8678
22.1147 | -105.8644
22.1185 | -105.8619
22.1223 | -105.8613
22.1261 | -105.8616
22.1298 | -105.8603
22.1336 | -105.8604
22.1374 | -105.8634
22.1412 | -105.8668
22.1449 | -105.8671
22.1487 | -105.8688
22.1525 | -105.8718
22.1563 | -105.8737
22.1601 | -105.8772
22.1638 | -105.8832
22.1676 | -105.8871
22.1714 | -105.8896
22.1752 | -105.8927
22.179 | -105.898
22.1827 | -105.9037
22.1865 | -105.9098
22.1903 | -105.9152
22.1941 | -105.9212
22.1979 | -105.9295
22.2016 | -105.9373
22.2054 | -105.9446
22.2092 | -105.9503
22.213 | -105.9566
22.2168 | -105.9649
22.2205 | -105.9746
22.2243 | -105.9834
22.2281 | -105.9921
22.2319 | -106.0012
22.2357 | -106.0117
22.2394 | -106.0224
22.2432 | -106.0335
22.247 | -106.0423
22.2508 | -106.0527
22.2546 | -106.063
22.2583 | -106.072
22.2621 | -106.0807
22.2659 | -106.0892
22.2697 | -106.0986
22.2735 | -106.1079
22.2772 | -106.116
22.281 | -106.1246
22.2848 | -106.1349
22.2886 | -106.143
22.2924 | -106.1491
22.2961 | -106.1548
22.2999 | -106.1609
22.3037 | -106.1692
22.3075 | -106.1769
22.3113 | -106.1832
22.315 | -106.1882
22.3188 | -106.1925
22.3226 | -106.1965
22.3264 | -106.1972
22.3302 | -106.1985
22.3339 | -106.1999
22.3377 | -106.2011
22.3415 | -106.2
22.3453 | -106.1991
22.349 | -106.1971
22.3528 | -106.1913
22.3566 | -106.1796
22.3604 | -106.1733
22.3642 | -106.1694
22.3679 | -106.1634
22.3717 | -106.1556
22.3755 | -106.149
22.3793 | -106.1415
22.3831 | -106.1334
22.3868 | -106.1238
22.3906 | -106.1144
22.3944 | -106.1067
22.3982 | -106.0986
22.402 | -106.0869
22.4057 | -106.075
22.4095 | -106.0637
22.4133 | -106.0525
22.4171 | -106.0422
22.4209 | -106.0329
22.4246 | -106.024
22.4284 | -106.016
22.4322 | -106.0084
22.436 | -105.9999
22.4398 | -105.9917
22.4435 | -105.9826
22.4473 | -105.9754
22.4511 | -105.9682
22.4549 | -105.9605
22.4587 | -105.9528
22.4624 | -105.9475
22.4662 | -105.9432
22.47 | -105.9388
22.4738 | -105.9337
22.4776 | -105.9268
22.4813 | -105.9193
22.4851 | -105.9139
22.4889 | -105.9117
22.4927 | -105.9092
22.4965 | -105.9058
22.5002 | -105.9023
22.504 | -105.8982
22.5078 | -105.8937
22.5116 | -105.89
22.5154 | -105.8854
22.5191 | -105.8799
22.5229 | -105.8747
22.5267 | -105.8713
22.5305 | -105.8681
22.5343 | -105.8626
22.538 | -105.8566
22.5418 | -105.8509
22.5456 | -105.8455
22.5494 | -105.8386
22.5532 | -105.8313
22.5569 | -105.8239
22.5607 | -105.8159
22.5645 | -105.8086
22.5683 | -105.7995
22.5721 | -105.7895
22.5758 | -105.7798
22.5796 | -105.7697
22.5834 | -105.7571
22.5872 | -105.7445
22.5909 | -105.7334
22.5947 | -105.7243
22.5985 | -105.7115
22.6023 | -105.6989
22.6061 | -105.6882
22.6098 | -105.6775
22.6136 | -105.6667
22.6174 | -105.6535
22.6212 | -105.64
22.625 | -105.6284
22.6287 | -105.6169
22.6325 | -105.6057
22.6363 | -105.5946
22.6401 | -105.5841
22.6439 | -105.5751
22.6476 | -105.566
22.6514 | -105.555
22.6552 | -105.5445
22.659 | -105.5348
22.6628 | -105.5245
22.6665 | -105.5143
22.6703 | -105.5043
22.6741 | -105.4951
22.6779 | -105.4869
22.6817 | -105.4784
22.6854 | -105.4698
22.6892 | -105.4618
22.693 | -105.4547
22.6968 | -105.4459
22.7006 | -105.4379
22.7043 | -105.4303
22.7081 | -105.4242
22.7119 | -105.4183
22.7157 | -105.411
22.7195 | -105.4045
22.7232 | -105.3986
22.727 | -105.3943
22.7308 | -105.3912
22.7346 | -105.3868
22.7384 | -105.3829
22.7421 | -105.3807
22.7459 | -105.3782
22.7497 | -105.3771
22.7535 | -105.3744
22.7573 | -105.3733
22.761 | -105.3735
22.7648 | -105.3748
22.7686 | -105.3758
22.7724 | -105.3765
22.7762 | -105.3788
22.7799 | -105.3811
22.7837 | -105.3846
22.7875 | -105.3899
22.7913 | -105.398
22.795 | -105.406
22.7988 | -105.4132
22.8026 | -105.4214
22.8064 | -105.4301
22.8102 | -105.4391
22.8139 | -105.4486
22.8177 | -105.4572
22.8215 | -105.4671
22.8253 | -105.4782
22.8291 | -105.4906
22.8328 | -105.5017
22.8366 | -105.5138
22.8404 | -105.526
22.8442 | -105.5363
22.848 | -105.5483
22.8517 | -105.5605
22.8555 | -105.573
22.8593 | -105.5848
22.8631 | -105.5938
22.8669 | -105.6022
22.8706 | -105.6136
22.8744 | -105.6237
22.8782 | -105.6308
22.882 | -105.6371
22.8858 | -105.6439
22.8895 | -105.6502
22.8933 | -105.655
22.8971 | -105.6571
22.9009 | -105.6537
22.9047 | -105.6577
22.9084 | -105.6603
22.9122 | -105.6606
22.916 | -105.6574
22.9198 | -105.6537
22.9236 | -105.6518
22.9273 | -105.6484
22.9311 | -105.642
22.9349 | -105.6344
22.9387 | -105.627
22.9425 | -105.6192
22.9462 | -105.6113
22.95 | -105.5993
22.9538 | -105.586
22.9576 | -105.5754
22.9614 | -105.5648
22.9651 | -105.5533
22.9689 | -105.5414
22.9727 | -105.5291
22.9765 | -105.5179
22.9803 | -105.5049
22.984 | -105.4918
22.9878 | -105.4778
22.9916 | -105.4641
22.9954 | -105.4526
22.9992 | -105.4396
23.0029 | -105.4256
23.0067 | -105.4119
23.0105 | -105.3989
23.0143 | -105.3863
23.018 | -105.3745
23.0218 | -105.3632
23.0256 | -105.3508
23.0294 | -105.3389
23.0332 | -105.3265
23.0369 | -105.3163
23.0407 | -105.3061
23.0445 | -105.2964
23.0483 | -105.2889
23.0521 | -105.2795
23.0558 | -105.2704
23.0596 | -105.2622
23.0634 | -105.2539
23.0672 | -105.2469
23.071 | -105.24
23.0747 | -105.2325
23.0785 | -105.2277
23.0823 | -105.2232
23.0861 | -105.2179
23.0899 | -105.2136
23.0936 | -105.2082
23.0974 | -105.2035
23.1012 | -105.2008
23.105 | -105.1985
23.1088 | -105.1941
23.1125 | -105.1883
23.1163 | -105.1823
23.1201 | -105.1783
23.1239 | -105.1776
23.1277 | -105.1766
23.1314 | -105.1747
23.1352 | -105.1721
23.139 | -105.1701
23.1428 | -105.1686
23.1466 | -105.1659
23.1503 | -105.164
23.1541 | -105.1628
23.1579 | -105.1618
23.1617 | -105.1612
23.1655 | -105.1609
23.1692 | -105.1609
23.173 | -105.1583
23.1768 | -105.1557
23.1806 | -105.1541
23.1844 | -105.1532
23.1881 | -105.1534
23.1919 | -105.1526
23.1957 | -105.1506
23.1995 | -105.1484
23.2033 | -105.146
23.207 | -105.1452
23.2108 | -105.1435
23.2146 | -105.1425
23.2184 | -105.1424
23.2221 | -105.1414
23.2259 | -105.1368
23.2297 | -105.1328
23.2335 | -105.1298
23.2373 | -105.1286
23.241 | -105.1266
23.2448 | -105.1264
23.2486 | -105.1252
23.2524 | -105.1243
23.2562 | -105.1223
23.2599 | -105.1222
23.2637 | -105.122
23.2675 | -105.12
23.2713 | -105.1205
23.2751 | -105.1216
23.2788 | -105.1238
23.2826 | -105.126
23.2864 | -105.1263
23.2902 | -105.1271
23.294 | -105.1271
23.2977 | -105.127
23.3015 | -105.1252
23.3053 | -105.1238
23.3091 | -105.1271
23.3129 | -105.1299
23.3166 | -105.1318
23.3204 | -105.1341
23.3242 | -105.1354
23.328 | -105.1359
23.3318 | -105.1376
23.3355 | -105.1404
23.3393 | -105.1404
23.3431 | -105.1426
23.3469 | -105.1445
23.3507 | -105.1461
23.3544 | -105.1494
23.3582 | -105.152
23.362 | -105.1545
23.3658 | -105.1571
23.3696 | -105.1597
23.3733 | -105.1592
23.3771 | -105.16
23.3809 | -105.1629
23.3847 | -105.1645
23.3885 | -105.1661
23.3922 | -105.167
23.396 | -105.1684
23.3998 | -105.1708
23.4036 | -105.1727
23.4074 | -105.1753
23.4111 | -105.1773
23.4149 | -105.18
23.4187 | -105.1837
23.4225 | -105.1875
23.4262 | -105.1914
23.43 | -105.1927
23.4338 | -105.1938
23.4376 | -105.1964
23.4414 | -105.1982
23.4451 | -105.1986
23.4489 | -105.2
23.4527 | -105.1965
23.4565 | -105.1936
23.4603 | -105.1935
23.464 | -105.1933
23.4678 | -105.192
23.4716 | -105.1909
23.4754 | -105.1907
23.4792 | -105.1902
23.4829 | -105.1881
23.4867 | -105.1865
23.4905 | -105.1871
23.4943 | -105.1879
23.4981 | -105.1881
23.5018 | -105.1894
23.5056 | -105.1916
23.5094 | -105.1927
23.5132 | -105.1946
23.517 | -105.1963
23.5207 | -105.1981
23.5245 | -105.2001
23.5283 | -105.2018
23.5321 | -105.2043
23.5359 | -105.2075
23.5396 | -105.2108
23.5434 | -105.2144
23.5472 | -105.2168
23.551 | -105.2204
23.5548 | -105.2276
23.5585 | -105.2357
23.5623 | -105.2435
23.5661 | -105.2498
23.5699 | -105.2565
23.5737 | -105.2651
23.5774 | -105.2738
23.5812 | -105.2816
23.585 | -105.2883
23.5888 | -105.296
23.5926 | -105.3051
23.5963 | -105.315
23.6001 | -105.3258
23.6039 | -105.3356
23.6077 | -105.3453
23.6115 | -105.3553
23.6152 | -105.3643
23.619 | -105.3739
23.6228 | -105.3837
23.6266 | -105.3946
23.6304 | -105.4052
23.6341 | -105.4167
23.6379 | -105.4261
23.6417 | -105.434
23.6455 | -105.4425
23.6492 | -105.4511
23.653 | -105.4598
23.6568 | -105.4697
23.6606 | -105.4793
23.6644 | -105.4886
23.6681 | -105.4968
23.6719 | -105.507
23.6757 | -105.5156
23.6795 | -105.5248
23.6833 | -105.5331
23.687 | -105.5395
23.6908 | -105.5463
23.6946 | -105.5539
23.6984 | -105.5601
23.7022 | -105.5683
23.7059 | -105.5784
23.7097 | -105.5847
23.7135 | -105.5889
23.7173 | -105.5943
23.7211 | -105.5992
23.7248 | -105.6048
23.7286 | -105.6081
23.7324 | -105.6098
23.7362 | -105.6125
23.74 | -105.6153
23.7437 | -105.6177
23.7475 | -105.6212
23.7513 | -105.6229
23.7551 | -105.6239
23.7589 | -105.6251
23.7626 | -105.6275
23.7664 | -105.6311
23.7702 | -105.6333
23.774 | -105.6365
23.7778 | -105.6406
23.7815 | -105.6427
23.7853 | -105.6455
23.7891 | -105.6536
23.7929 | -105.6598
23.7967 | -105.6618
23.8004 | -105.6636
23.8042 | -105.668
23.808 | -105.6709
23.8118 | -105.673
23.8156 | -105.6763
23.8193 | -105.6801
23.8231 | -105.6815
23.8269 | -105.6821
23.8307 | -105.6833
23.8345 | -105.6859
23.8382 | -105.6883
23.842 | -105.6903
23.8458 | -105.6926
23.8496 | -105.6933
23.8533 | -105.6968
23.8571 | -105.7014
23.8609 | -105.7055
23.8647 | -105.7084
23.8685 | -105.7131
23.8722 | -105.7186
23.876 | -105.7222
23.8798 | -105.7256
23.8836 | -105.7285
23.8874 | -105.7315
23.8911 | -105.7332
23.8949 | -105.737
23.8987 | -105.7411
23.9025 | -105.7457
23.9063 | -105.7485
23.91 | -105.7495
23.9138 | -105.7516
23.9176 | -105.7557
23.9214 | -105.7615
23.9252 | -105.767
23.9289 | -105.7699
23.9327 | -105.7715
23.9365 | -105.7736
23.9403 | -105.776
23.9441 | -105.781
23.9478 | -105.786
23.9516 | -105.79
23.9554 | -105.7947
23.9592 | -105.8001
23.963 | -105.8043
23.9667 | -105.8086
23.9705 | -105.8136
23.9743 | -105.8199
23.9781 | -105.8248
23.9819 | -105.8289
23.9856 | -105.835
23.9894 | -105.8428
23.9932 | -105.8479
23.997 | -105.8512
24.0008 | -105.8561
24.0045 | -105.8624
24.0083 | -105.8685
24.0121 | -105.8732
24.0159 | -105.8779
24.0197 | -105.884
24.0234 | -105.8922
24.0272 | -105.9015
24.031 | -105.9108
24.0348 | -105.9203
24.0386 | -105.9291
24.0423 | -105.9362
24.0461 | -105.9451
24.0499 | -105.9546
24.0537 | -105.9623
24.0575 | -105.9706
24.0612 | -105.9808
24.065 | -105.9914
24.0688 | -106.0007
24.0726 | -106.0094
24.0763 | -106.0165
24.0801 | -106.0239
24.0839 | -106.033
24.0877 | -106.04
24.0915 | -106.0476
24.0952 | -106.0577
24.099 | -106.0665
24.1028 | -106.0738
24.1066 | -106.0809
24.1104 | -106.0882
24.1141 | -106.0921
24.1179 | -106.0943
24.1217 | -106.0991
24.1255 | -106.1041
24.1293 | -106.1074
24.133 | -106.1092
24.1368 | -106.112
24.1406 | -106.1156
24.1444 | -106.117
24.1482 | -106.1188
24.1519 | -106.1209
24.1557 | -106.1196
24.1595 | -106.1174
24.1633 | -106.1171
24.1671 | -106.1167
24.1708 | -106.1144
24.1746 | -106.1123
24.1784 | -106.108
24.1822 | -106.1013
24.186 | -106.0959
24.1897 | -106.0902
24.1935 | -106.0825
24.1973 | -106.0735
24.2011 | -106.0646
24.2049 | -106.0563
24.2086 | -106.0457
24.2124 | -106.0365
24.2162 | -106.0278
24.22 | -106.0196
24.2238 | -106.0121
24.2275 | -106.0037
24.2313 | -105.9949
24.2351 | -105.9844
24.2389 | -105.9754
24.2427 | -105.9654
24.2464 | -105.9558
24.2502 | -105.9486
24.254 | -105.94
24.2578 | -105.9301
24.2616 | -105.9215
24.2653 | -105.9132
24.2691 | -105.9038
24.2729 | -105.8928
24.2767 | -105.8794
24.2804 | -105.8613
24.2842 | -105.8362
24.288 | -105.8121
24.2918 | -105.7891
24.2956 | -105.7601
24.2993 | -105.7285
24.3031 | -105.7017
24.3069 | -105.6747
24.3107 | -105.6435
24.3145 | -105.6095
24.3182 | -105.5748
24.322 | -105.5412
24.3258 | -105.5071
24.3296 | -105.472
24.3334 | -105.4346
24.3371 | -105.398
24.3409 | -105.3601
24.3447 | -105.3197
24.3485 | -105.2781
24.3523 | -105.2353
24.356 | -105.1922
24.3598 | -105.1486
24.3636 | -105.1021
24.3674 | -105.0542
24.3712 | -105.0061
24.3749 | -104.9581
24.3787 | -104.9088
24.3825 | -104.8589
24.3863 | -104.8094
24.3901 | -104.7601
24.3938 | -104.7125
24.3976 | -104.6668
24.4014 | -104.6204
24.4052 | -104.574
24.409 | -104.5281
24.4127 | -104.485
24.4165 | -104.4454
24.4203 | -104.4078
24.4241 | -104.3709
24.4279 | -104.336
24.4316 | -104.3037
24.4354 | -104.2747
24.4392 | -104.2477
24.443 | -104.2241
24.4468 | -104.2039
24.4505 | -104.1871
24.4543 | -104.1746
24.4581 | -104.1673
24.4619 | -104.1636
24.4657 | -104.1619
24.4694 | -104.1625
24.4732 | -104.1662
24.477 | -104.1735
24.4808 | -104.1851
24.4846 | -104.2009
24.4883 | -104.2188
24.4921 | -104.239
24.4959 | -104.2604
24.4997 | -104.2828
24.5034 | -104.3086
24.5072 | -104.3387
24.511 | -104.3711
24.5148 | -104.4054
24.5186 | -104.4402
24.5223 | -104.4772
24.5261 | -104.5149
24.5299 | -104.5517
24.5337 | -104.5892
24.5375 | -104.6244
24.5412 | -104.6609
24.545 | -104.6978
24.5488 | -104.7342
24.5526 | -104.7695
24.5564 | -104.8039
24.5601 | -104.8351
24.5639 | -104.8602
24.5677 | -104.8761
24.5715 | -104.8778
24.5753 | -104.8594
24.579 | -104.8278
24.5828 | -104.7695
24.5866 | -104.7124
24.5904 | -104.6456
24.5942 | -104.5706
24.5979 | -104.4972
24.6017 | -104.4223
24.6055 | -104.3511
24.6093 | -104.2757
24.6131 | -104.1951
24.6168 | -104.1161
24.6206 | -104.0326
24.6244 | -103.9436
24.6282 | -103.8522
24.632 | -103.7574
24.6357 | -103.6569
24.6395 | -103.5526
24.6433 | -103.4461
24.6471 | -103.3356
24.6509 | -103.2196
24.6546 | -103.1011
24.6584 | -102.9834
24.6622 | -102.8656
24.666 | -102.7467
24.6698 | -102.6273
24.6735 | -102.5087
24.6773 | -102.3918
24.6811 | -102.2775
24.6849 | -102.1652
24.6887 | -102.0564
24.6924 | -101.9515
24.6962 | -101.8497
24.7 | -101.7519
24.7038 | -101.6597
24.7076 | -101.572
24.7113 | -101.4895
24.7151 | -101.4133
24.7189 | -101.3424
24.7227 | -101.2766
24.7264 | -101.2172
24.7302 | -101.1645
24.734 | -101.1186
24.7378 | -101.0791
24.7416 | -101.046
24.7453 | -101.0197
24.7491 | -101.0002
24.7529 | -100.9879
24.7567 | -100.9826
24.7605 | -100.9844
24.7642 | -100.9928
24.768 | -101.0085
24.7718 | -101.0314
24.7756 | -101.0613
24.7794 | -101.098
24.7831 | -101.1415
24.7869 | -101.1921
24.7907 | -101.2497
24.7945 | -101.3139
24.7983 | -101.3842
24.802 | -101.4603
24.8058 | -101.5429
24.8096 | -101.6313
24.8134 | -101.725
24.8172 | -101.824
24.8209 | -101.9283
24.8247 | -102.0376
24.8285 | -102.1504
24.8323 | -102.266
24.8361 | -102.3865
24.8398 | -102.5111
24.8436 | -102.6389
24.8474 | -102.7701
24.8512 | -102.9013
24.855 | -103.0321
24.8587 | -103.1635
24.8625 | -103.2951
24.8663 | -103.4269
24.8701 | -103.5574
24.8739 | -103.6878
24.8776 | -103.8164
24.8814 | -103.9419
24.8852 | -104.0631
24.889 | -104.1815
24.8928 | -104.2955
24.8965 | -104.4049
24.9003 | -104.5109
24.9041 | -104.612
24.9079 | -104.7089
24.9117 | -104.8021
24.9154 | -104.8927
24.9192 | -104.9842
24.923 | -105.08
24.9268 | -105.1822
24.9305 | -105.282
24.9343 | -105.3929
24.9381 | -105.4821
24.9419 | -105.5656
24.9457 | -105.6398
24.9494 | -105.6914
24.9532 | -105.7286
24.957 | -105.7474
24.9608 | -105.7596
24.9646 | -105.768
24.9683 | -105.7664
24.9721 | -105.7644
24.9759 | -105.7646
24.9797 | -105.7634
24.9835 | -105.7626
24.9872 | -105.7624
24.991 | -105.7608
24.9948 | -105.758
24.9986 | -105.7562
25.0024 | -105.7552
25.0061 | -105.7546
25.0099 | -105.7536
25.0137 | -105.7527
25.0175 | -105.7513
25.0213 | -105.7478
25.025 | -105.7443
25.0288 | -105.7407
25.0326 | -105.738
25.0364 | -105.7358
25.0402 | -105.7322
25.0439 | -105.7281
25.0477 | -105.7256
25.0515 | -105.7251
25.0553 | -105.7212
25.0591 | -105.7176
25.0628 | -105.7153
25.0666 | -105.7115
25.0704 | -105.7089
25.0742 | -105.7065
25.078 | -105.7036
25.0817 | -105.6981
25.0855 | -105.6935
25.0893 | -105.6904
25.0931 | -105.6879
25.0969 | -105.6847
25.1006 | -105.6802
25.1044 | -105.6764
25.1082 | -105.6731
25.112 | -105.6692
25.1158 | -105.6663
25.1195 | -105.6649
25.1233 | -105.663
25.1271 | -105.6616
25.1309 | -105.6603
25.1347 | -105.6588
25.1384 | -105.6559
25.1422 | -105.6556
25.146 | -105.6549
25.1498 | -105.6538
25.1535 | -105.6518
25.1573 | -105.6532
25.1611 | -105.6537
25.1649 | -105.6539
25.1687 | -105.6542
25.1724 | -105.6564
25.1762 | -105.6592
25.18 | -105.6608
25.1838 | -105.6636
25.1876 | -105.6672
25.1913 | -105.6694
25.1951 | -105.6696
25.1989 | -105.67
25.2027 | -105.6709
25.2065 | -105.6732
25.2102 | -105.6748
25.214 | -105.6742
25.2178 | -105.6729
25.2216 | -105.6723
25.2254 | -105.6736
25.2291 | -105.6735
25.2329 | -105.6722
25.2367 | -105.6709
25.2405 | -105.6676
25.2443 | -105.6627
25.248 | -105.6608
25.2518 | -105.6578
25.2556 | -105.6524
25.2594 | -105.6462
25.2632 | -105.6394
25.2669 | -105.6311
25.2707 | -105.6236
25.2745 | -105.6152
25.2783 | -105.6074
25.2821 | -105.5998
25.2858 | -105.5941
25.2896 | -105.5845
25.2934 | -105.5727
25.2972 | -105.5618
25.301 | -105.5497
25.3047 | -105.5373
25.3085 | -105.5254
25.3123 | -105.5129
25.3161 | -105.5003
25.3199 | -105.4882
25.3236 | -105.4756
25.3274 | -105.4628
25.3312 | -105.4486
25.335 | -105.4361
25.3388 | -105.4225
25.3425 | -105.4062
25.3463 | -105.3912
25.3501 | -105.3777
25.3539 | -105.3644
25.3576 | -105.3498
25.3614 | -105.3352
25.3652 | -105.3204
25.369 | -105.3054
25.3728 | -105.2924
25.3765 | -105.2801
25.3803 | -105.2682
25.3841 | -105.2565
25.3879 | -105.2448
25.3917 | -105.2336
25.3954 | -105.2245
25.3992 | -105.2152
25.403 | -105.2056
25.4068 | -105.195
25.4106 | -105.1859
25.4143 | -105.1765
25.4181 | -105.1672
25.4219 | -105.1566
25.4257 | -105.1476
25.4295 | -105.1423
25.4332 | -105.136
25.437 | -105.1288
25.4408 | -105.1244
25.4446 | -105.12
25.4484 | -105.1152
25.4521 | -105.1093
25.4559 | -105.1042
25.4597 | -105.1016
25.4635 | -105.0986
25.4673 | -105.0942
25.471 | -105.0883
25.4748 | -105.0818
25.4786 | -105.0755
25.4824 | -105.0709
25.4862 | -105.0665
25.4899 | -105.0613
25.4937 | -105.0579
25.4975 | -105.0552
25.5013 | -105.0494
25.5051 | -105.0436
25.5088 | -105.0392
25.5126 | -105.0363
25.5164 | -105.0311
25.5202 | -105.0261
25.524 | -105.0227
25.5277 | -105.0211
25.5315 | -105.0189
25.5353 | -105.0161
25.5391 | -105.0138
25.5429 | -105.0107
25.5466 | -105.0073
25.5504 | -105.0054
25.5542 | -105.0042
25.558 | -105.003
25.5617 | -105.001
25.5655 | -104.9996
25.5693 | -104.9978
25.5731 | -104.9967
25.5769 | -104.9972
25.5806 | -104.9984
25.5844 | -104.9981
25.5882 | -104.9983
25.592 | -104.999
25.5958 | -104.9991
25.5995 | -105.0005
25.6033 | -105.0028
25.6071 | -105.0054
25.6109 | -105.0095
25.6147 | -105.0133
25.6184 | -105.0169
25.6222 | -105.0193
25.626 | -105.0225
25.6298 | -105.0266
25.6336 | -105.0313
25.6373 | -105.036
25.6411 | -105.039
25.6449 | -105.0419
25.6487 | -105.0429
25.6525 | -105.0457
25.6562 | -105.0491
25.66 | -105.0515
25.6638 | -105.0561
25.6676 | -105.0604
25.6714 | -105.0638
25.6751 | -105.0676
25.6789 | -105.0705
25.6827 | -105.0725
25.6865 | -105.0759
25.6903 | -105.0802
25.694 | -105.086
25.6978 | -105.0894
25.7016 | -105.0942
25.7054 | -105.1008
25.7092 | -105.1064
25.7129 | -105.1111
25.7167 | -105.1138
25.7205 | -105.1178
25.7243 | -105.1223
25.7281 | -105.1265
25.7318 | -105.1303
25.7356 | -105.1338
25.7394 | -105.1389
25.7432 | -105.1424
25.747 | -105.1466
25.7507 | -105.1512
25.7545 | -105.1558
25.7583 | -105.1597
25.7621 | -105.1645
25.7659 | -105.1707
25.7696 | -105.1779
25.7734 | -105.184
25.7772 | -105.1913
25.781 | -105.1995
25.7847 | -105.2088
25.7885 | -105.2164
25.7923 | -105.2218
25.7961 | -105.2299
25.7999 | -105.2388
25.8036 | -105.2467
25.8074 | -105.2548
25.8112 | -105.263
25.815 | -105.2732
25.8188 | -105.2829
25.8225 | -105.2919
25.8263 | -105.3025
25.8301 | -105.3147
25.8339 | -105.3281
25.8377 | -105.3405
25.8414 | -105.3517
25.8452 | -105.3606
25.849 | -105.3705
25.8528 | -105.3825
25.8566 | -105.3936
25.8603 | -105.4038
25.8641 | -105.4139
25.8679 | -105.425
25.8717 | -105.4347
25.8755 | -105.4437
25.8792 | -105.4533
25.883 | -105.4624
25.8868 | -105.4701
25.8906 | -105.4769
25.8944 | -105.4857
25.8981 | -105.4942
25.9019 | -105.5021
25.9057 | -105.5106
25.9095 | -105.5181
25.9133 | -105.5242
25.917 | -105.5314
25.9208 | -105.5381
25.9246 | -105.5449
25.9284 | -105.5501
25.9322 | -105.5548
25.9359 | -105.5603
25.9397 | -105.5644
25.9435 | -105.5688
25.9473 | -105.5723
25.9511 | -105.5756
25.9548 | -105.5802
25.9586 | -105.5848
25.9624 | -105.588
25.9662 | -105.5904
25.97 | -105.592
25.9737 | -105.5935
25.9775 | -105.5963
25.9813 | -105.5975
25.9851 | -105.5992
25.9888 | -105.6024
25.9926 | -105.6045
25.9964 | -105.6046
26.0002 | -105.6046
26.004 | -105.606
26.0077 | -105.6077
26.0115 | -105.6099
26.0153 | -105.6137
26.0191 | -105.6178
26.0229 | -105.621
26.0266 | -105.6225
26.0304 | -105.6237
26.0342 | -105.6254
26.038 | -105.6273
26.0418 | -105.6283
26.0455 | -105.6314
26.0493 | -105.637
26.0531 | -105.641
26.0569 | -105.6451
26.0607 | -105.6477
26.0644 | -105.6512
26.0682 | -105.6554
26.072 | -105.6608
26.0758 | -105.6673
26.0796 | -105.6736
26.0833 | -105.6798
26.0871 | -105.6875
26.0909 | -105.6953
26.0947 | -105.7041
26.0985 | -105.7132
26.1022 | -105.7217
26.106 | -105.7284
26.1098 | -105.7362
26.1136 | -105.7444
26.1174 | -105.7522
26.1211 | -105.7615
26.1249 | -105.7719
26.1287 | -105.7808
26.1325 | -105.7877
26.1363 | -105.7964
26.14 | -105.806
26.1438 | -105.8152
26.1476 | -105.8243
26.1514 | -105.8328
26.1552 | -105.8399
26.1589 | -105.847
26.1627 | -105.8537
26.1665 | -105.8609
26.1703 | -105.8644
26.1741 | -105.8712
26.1778 | -105.8775
26.1816 | -105.8822
26.1854 | -105.8862
26.1892 | -105.8895
26.193 | -105.8937
26.1967 | -105.8973
26.2005 | -105.8998
26.2043 | -105.9023
26.2081 | -105.9044
26.2118 | -105.9058
26.2156 | -105.9097
26.2194 | -105.9123
26.2232 | -105.9123
26.227 | -105.9129
26.2307 | -105.914
26.2345 | -105.9154
26.2383 | -105.915
26.2421 | -105.9151
26.2459 | -105.916
26.2496 | -105.9171
26.2534 | -105.9173
26.2572 | -105.9169
26.261 | -105.9168
26.2648 | -105.914
26.2685 | -105.9128
26.2723 | -105.9122
26.2761 | -105.9108
26.2799 | -105.9115
26.2837 | -105.9135
26.2874 | -105.9149
26.2912 | -105.9171
26.295 | -105.918
26.2988 | -105.9214
26.3026 | -105.9252
26.3063 | -105.9282
26.3101 | -105.9314
26.3139 | -105.9327
26.3177 | -105.9347
26.3215 | -105.9381
26.3252 | -105.9405
26.329 | -105.9429
26.3328 | -105.9466
26.3366 | -105.949
26.3404 | -105.9509
26.3441 | -105.9521
26.3479 | -105.9536
26.3517 | -105.9581
26.3555 | -105.9629
26.3593 | -105.9674
26.363 | -105.9733
26.3668 | -105.9785
26.3706 | -105.9828
26.3744 | -105.987
26.3782 | -105.9931
26.3819 | -105.998
26.3857 | -106.0023
26.3895 | -106.0081
26.3933 | -106.0141
26.3971 | -106.021
26.4008 | -106.0279
26.4046 | -106.0351
26.4084 | -106.0416
26.4122 | -106.0486
26.4159 | -106.0536
26.4197 | -106.058
26.4235 | -106.0652
26.4273 | -106.0713
26.4311 | -106.0776
26.4348 | -106.0825
26.4386 | -106.0858
26.4424 | -106.089
26.4462 | -106.0917
26.45 | -106.0959
26.4537 | -106.1016
26.4575 | -106.1046
26.4613 | -106.1064
26.4651 | -106.1071
26.4689 | -106.1101
26.4726 | -106.1132
26.4764 | -106.1167
26.4802 | -106.1198
26.484 | -106.1207
26.4878 | -106.1207
26.4915 | -106.1196
26.4953 | -106.1185
26.4991 | -106.1188
26.5029 | -106.1184
26.5067 | -106.1162
26.5104 | -106.1141
26.5142 | -106.112
26.518 | -106.1075
26.5218 | -106.1024
26.5256 | -106.0984
26.5293 | -106.094
26.5331 | -106.0896
26.5369 | -106.0843
26.5407 | -106.0777
26.5445 | -106.07
26.5482 | -106.0618
26.552 | -106.0549
26.5558 | -106.0482
26.5596 | -106.0393
26.5634 | -106.0309
26.5671 | -106.0214
26.5709 | -106.0119
26.5747 | -106.003
26.5785 | -105.9943
26.5823 | -105.9862
26.586 | -105.9789
26.5898 | -105.9691
26.5936 | -105.9595
26.5974 | -105.9521
26.6012 | -105.9421
26.6049 | -105.9321
26.6087 | -105.9236
26.6125 | -105.9144
26.6163 | -105.9029
26.6201 | -105.8927
26.6238 | -105.882
26.6276 | -105.8725
26.6314 | -105.8635
26.6352 | -105.8538
26.6389 | -105.8446
26.6427 | -105.8354
26.6465 | -105.825
26.6503 | -105.815
26.6541 | -105.8055
26.6578 | -105.7957
26.6616 | -105.7871
26.6654 | -105.7786
26.6692 | -105.7711
26.673 | -105.7625
26.6767 | -105.7536
26.6805 | -105.7454
26.6843 | -105.7371
26.6881 | -105.7311
26.6919 | -105.7276
26.6956 | -105.722
26.6994 | -105.717
26.7032 | -105.7135
26.707 | -105.71
26.7108 | -105.7076
26.7145 | -105.7058
26.7183 | -105.7067
26.7221 | -105.7059
26.7259 | -105.7051
26.7297 | -105.7044
26.7334 | -105.7036
26.7372 | -105.7023
26.741 | -105.7006
26.7448 | -105.7015
26.7486 | -105.7034
26.7523 | -105.7053
26.7561 | -105.7065
26.7599 | -105.7091
26.7637 | -105.7125
26.7675 | -105.7148
26.7712 | -105.7188
26.775 | -105.7215
26.7788 | -105.7257
26.7826 | -105.7314
26.7864 | -105.7375
26.7901 | -105.7427
26.7939 | -105.7492
26.7977 | -105.7554
26.8015 | -105.7605
26.8053 | -105.767
26.809 | -105.7744
26.8128 | -105.784
26.8166 | -105.7933
26.8204 | -105.8016
26.8242 | -105.8113
26.8279 | -105.8201
26.8317 | -105.8291
26.8355 | -105.8392
26.8393 | -105.8489
26.8431 | -105.8584
26.8468 | -105.867
26.8506 | -105.8762
26.8544 | -105.885
26.8582 | -105.8929
26.8619 | -105.9027
26.8657 | -105.9129
26.8695 | -105.921
26.8733 | -105.9281
26.8771 | -105.9356
26.8808 | -105.9445
26.8846 | -105.9527
26.8884 | -105.9607
26.8922 | -105.9687
26.896 | -105.9772
26.8997 | -105.9856
26.9035 | -105.9936
26.9073 | -106.0004
26.9111 | -106.0044
26.9149 | -106.0075
26.9186 | -106.0124
26.9224 | -106.0166
26.9262 | -106.0206
26.93 | -106.023
26.9338 | -106.0252
26.9375 | -106.0289
26.9413 | -106.0327
26.9451 | -106.0345
26.9489 | -106.0356
26.9527 | -106.0375
26.9564 | -106.0385
26.9602 | -106.0379
26.964 | -106.037
26.9678 | -106.0371
26.9716 | -106.038
26.9753 | -106.0369
26.9791 | -106.0369
26.9829 | -106.0377
26.9867 | -106.0375
26.9905 | -106.0386
26.9942 | -106.0418
26.998 | -106.0417
27.0018 | -106.0397
27.0056 | -106.0372
27.0094 | -106.036
27.0131 | -106.034
27.0169 | -106.0338
27.0207 | -106.0317
27.0245 | -106.0307
27.0283 | -106.0314
27.032 | -106.0315
27.0358 | -106.0315
27.0396 | -106.0311
27.0434 | -106.0315
27.0472 | -106.0321
27.0509 | -106.0315
27.0547 | -106.0328
27.0585 | -106.0347
27.0623 | -106.037
27.066 | -106.0389
27.0698 | -106.0398
27.0736 | -106.042
27.0774 | -106.0442
27.0812 | -106.0464
27.0849 | -106.0501
27.0887 | -106.0542
27.0925 | -106.0578
27.0963 | -106.062
27.1001 | -106.0672
27.1038 | -106.0725
27.1076 | -106.0794
27.1114 | -106.0863
27.1152 | -106.0918
27.119 | -106.0979
27.1227 | -106.1034
27.1265 | -106.1105
27.1303 | -106.1188
27.1341 | -106.1267
27.1379 | -106.1358
27.1416 | -106.1437
27.1454 | -106.1525
27.1492 | -106.1605
27.153 | -106.1692
27.1568 | -106.179
27.1605 | -106.1879
27.1643 | -106.1974
27.1681 | -106.2085
27.1719 | -106.2208
27.1757 | -106.2305
27.1794 | -106.2401
27.1832 | -106.2498
27.187 | -106.2594
27.1908 | -106.2693
27.1946 | -106.2779
27.1983 | -106.2864
27.2021 | -106.2959
27.2059 | -106.3033
27.2097 | -106.3112
27.2135 | -106.3186
27.2172 | -106.326
27.221 | -106.3322
27.2248 | -106.3381
27.2286 | -106.343
27.2324 | -106.3478
27.2361 | -106.3545
27.2399 | -106.3606
27.2437 | -106.3663
27.2475 | -106.3694
27.2513 | -106.3744
27.255 | -106.3798
27.2588 | -106.3836
27.2626 | -106.382
27.2664 | -106.3837
27.2702 | -106.3862
27.2739 | -106.389
27.2777 | -106.391
27.2815 | -106.3928
27.2853 | -106.393
27.289 | -106.3923
27.2928 | -106.393
27.2966 | -106.393
27.3004 | -106.3934
27.3042 | -106.3922
27.3079 | -106.392
27.3117 | -106.3898
27.3155 | -106.3865
27.3193 | -106.3825
27.3231 | -106.3779
27.3268 | -106.3732
27.3306 | -106.3689
27.3344 | -106.3654
27.3382 | -106.3611
27.342 | -106.3575
27.3457 | -106.355
27.3495 | -106.3509
27.3533 | -106.3453
27.3571 | -106.3377
27.3609 | -106.3327
27.3646 | -106.3258
27.3684 | -106.3186
27.3722 | -106.3123
27.376 | -106.3051
27.3798 | -106.2973
27.3835 | -106.291
27.3873 | -106.287
27.3911 | -106.2839
27.3949 | -106.2792
27.3987 | -106.2743
27.4024 | -106.2668
27.4062 | -106.2605
27.41 | -106.2547
27.4138 | -106.2469
27.4176 | -106.2392
27.4213 | -106.2316
27.4251 | -106.2236
27.4289 | -106.2155
27.4327 | -106.2086
27.4365 | -106.2035
27.4402 | -106.1971
27.444 | -106.1893
27.4478 | -106.1824
27.4516 | -106.1752
27.4554 | -106.1685
27.4591 | -106.1639
27.4629 | -106.1586
27.4667 | -106.1516
27.4705 | -106.1444
27.4743 | -106.1387
27.478 | -106.1332
27.4818 | -106.1288
27.4856 | -106.1242
27.4894 | -106.1183
27.4931 | -106.1117
27.4969 | -106.107
27.5007 | -106.1054
27.5045 | -106.1027
27.5083 | -106.0987
27.512 | -106.0932
27.5158 | -106.0869
27.5196 | -106.0798
27.5234 | -106.0744
27.5272 | -106.0706
27.5309 | -106.0653
27.5347 | -106.0612
27.5385 | -106.0558
27.5423 | -106.0514
27.5461 | -106.045
27.5498 | -106.0364
27.5536 | -106.0269
27.5574 | -106.0189
27.5612 | -106.0135
27.565 | -106.0081
27.5687 | -106.0017
27.5725 | -105.9952
27.5763 | -105.9884
27.5801 | -105.982
27.5839 | -105.9757
27.5876 | -105.9675
27.5914 | -105.9601
27.5952 | -105.9534
27.599 | -105.9463
27.6028 | -105.9382
27.6065 | -105.9324
27.6103 | -105.9273
27.6141 | -105.922
27.6179 | -105.9157
27.6217 | -105.9086
27.6254 | -105.9024
27.6292 | -105.8964
27.633 | -105.8906
27.6368 | -105.8817
27.6406 | -105.8716
27.6443 | -105.8622
27.6481 | -105.8524
27.6519 | -105.8403
27.6557 | -105.8235
27.6595 | -105.8032
27.6632 | -105.7811
27.667 | -105.7492
27.6708 | -105.7094
27.6746 | -105.6675
27.6784 | -105.6153
27.6821 | -105.5529
27.6859 | -105.4824
27.6897 | -105.3988
27.6935 | -105.3061
27.6973 | -105.2113
27.701 | -105.115
27.7048 | -105.016
27.7086 | -104.9155
27.7124 | -104.813
27.7161 | -104.7039
27.7199 | -104.5872
27.7237 | -104.459
27.7275 | -104.3225
27.7313 | -104.179
27.735 | -104.036
27.7388 | -103.8937
27.7426 | -103.7493
27.7464 | -103.6046
27.7502 | -103.4546
27.7539 | -103.2998
27.7577 | -103.1438
27.7615 | -102.9871
27.7653 | -102.8289
27.7691 | -102.6693
27.7728 | -102.5083
27.7766 | -102.3482
27.7804 | -102.1872
27.7842 | -102.0241
27.788 | -101.8631
27.7917 | -101.7047
27.7955 | -101.5469
27.7993 | -101.3894
27.8031 | -101.2334
27.8069 | -101.0527
27.8106 | -100.9051
27.8144 | -100.76
27.8182 | -100.6173
27.822 | -100.4766
27.8258 | -100.3387
27.8295 | -100.2017
27.8333 | -100.0677
27.8371 | -99.9361
27.8409 | -99.8061
27.8447 | -99.6818
27.8484 | -99.562
27.8522 | -99.4453
27.856 | -99.3291
27.8598 | -99.2163
27.8636 | -99.1034
27.8673 | -98.9928
27.8711 | -98.8837
27.8749 | -98.7776
27.8787 | -98.6657
27.8825 | -98.5588
27.8862 | -98.4539
27.89 | -98.3497
27.8938 | -98.2454
27.8976 | -98.1416
27.9014 | -98.0357
27.9051 | -97.9293
27.9089 | -97.827
27.9127 | -97.7276
27.9165 | -97.6276
27.9202 | -97.5285
27.924 | -97.432
27.9278 | -97.3343
27.9316 | -97.2387
27.9354 | -97.1439
27.9391 | -97.0426
27.9429 | -96.9421
27.9467 | -96.8432
27.9505 | -96.7441
27.9543 | -96.6464
27.958 | -96.547
27.9618 | -96.4505
27.9656 | -96.3536
27.9694 | -96.256
27.9732 | -96.1598
27.9769 | -96.0631
27.9807 | -95.9669
27.9845 | -95.8705
27.9883 | -95.7736
27.9921 | -95.6769
27.9958 | -95.5808
27.9996 | -95.4858
28.0034 | -95.3916