## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 228 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 222
Hood | 02-92B/P | 161
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1009 | hPa |  | PRC PRO 01
Humidity | 41.9 | %/Hr
Internal temperature of calibration bench | 22.59 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1014 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.9 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.87 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.96 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3108 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 07/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 25/05/2018
Engineer | Anthony HUE | Approval |  | 29/05/2018
CEO | Robert GLEMOT | Approval |  | 29/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5755 | 57.517
0.012209 | -36.1546 | 38.3594
0.018313 | -35.0934 | 28.1955
0.024417 | -34.633 | 21.9765
0.030521 | -34.5181 | 18.2
0.036626 | -34.3729 | 15.1491
0.04273 | -34.2777 | 12.7931
0.048834 | -34.2321 | 11.2717
0.054938 | -34.1678 | 10.1279
0.061043 | -34.1073 | 9.04
0.067147 | -34.0956 | 8.2223
0.073251 | -34.0958 | 7.4664
0.079355 | -34.1142 | 6.809
0.08546 | -34.1288 | 6.2204
0.091564 | -34.1241 | 5.7833
0.097668 | -34.1053 | 5.5313
0.10377 | -34.0827 | 5.2811
0.10988 | -34.065 | 5.0038
0.11598 | -34.0555 | 4.713
0.12209 | -34.0571 | 4.4311
0.12819 | -34.067 | 4.1486
0.13429 | -34.0782 | 3.8613
0.1404 | -34.0817 | 3.5857
0.1465 | -34.0838 | 3.3306
0.15261 | -34.0812 | 3.1132
0.15871 | -34.0746 | 2.9533
0.16482 | -34.0663 | 2.8518
0.17092 | -34.0638 | 2.8118
0.17702 | -34.0614 | 2.7395
0.18313 | -34.0604 | 2.6608
0.18923 | -34.06 | 2.5517
0.19534 | -34.0588 | 2.425
0.20144 | -34.0575 | 2.2834
0.20754 | -34.0544 | 2.1635
0.21365 | -34.0533 | 2.0494
0.21975 | -34.0526 | 1.9521
0.22586 | -34.0534 | 1.8651
0.23196 | -34.0555 | 1.793
0.23807 | -34.0593 | 1.7303
0.24417 | -34.0632 | 1.6719
0.25027 | -34.0669 | 1.6099
0.25638 | -34.0692 | 1.5486
0.26248 | -34.0709 | 1.4827
0.26859 | -34.0718 | 1.4182
0.27469 | -34.0721 | 1.348
0.2808 | -34.0732 | 1.2749
0.2869 | -34.074 | 1.2044
0.293 | -34.075 | 1.1376
0.29911 | -34.0759 | 1.0774
0.30521 | -34.077 | 1.0182
0.31132 | -34.0778 | 0.97913
0.31742 | -34.079 | 0.93602
0.32353 | -34.0794 | 0.90094
0.32963 | -34.0803 | 0.87194
0.33573 | -34.0814 | 0.85178
0.34184 | -34.0827 | 0.83087
0.34794 | -34.0848 | 0.80828
0.35405 | -34.0869 | 0.785
0.36015 | -34.0894 | 0.7606
0.36626 | -34.0913 | 0.73417
0.37236 | -34.0933 | 0.69859
0.37846 | -34.0954 | 0.6555
0.38457 | -34.0961 | 0.59913
0.39067 | -34.0971 | 0.55325
0.39678 | -34.0976 | 0.50211
0.40288 | -34.098 | 0.4487
0.40899 | -34.0978 | 0.39725
0.41509 | -34.097 | 0.34674
0.42119 | -34.0969 | 0.30352
0.4273 | -34.0968 | 0.2575
0.4334 | -34.0966 | 0.2116
0.43951 | -34.0969 | 0.16838
0.44561 | -34.0977 | 0.12875
0.45172 | -34.0992 | 0.097239
0.45782 | -34.1 | 0.060368
0.46392 | -34.1012 | 0.024448
0.47003 | -34.1022 | -0.0076823
0.47613 | -34.103 | -0.041389
0.48224 | -34.1035 | -0.075784
0.48834 | -34.1026 | -0.10993
0.49445 | -34.1024 | -0.14457
0.50055 | -34.1018 | -0.17612
0.50665 | -34.101 | -0.20979
0.51276 | -34.1 | -0.24316
0.51886 | -34.099 | -0.27646
0.52497 | -34.0977 | -0.30015
0.53107 | -34.0969 | -0.32913
0.53717 | -34.0962 | -0.35688
0.54328 | -34.0953 | -0.37933
0.54938 | -34.0951 | -0.40154
0.55549 | -34.0951 | -0.42214
0.56159 | -34.0961 | -0.4338
0.5677 | -34.0964 | -0.44837
0.5738 | -34.0975 | -0.46458
0.5799 | -34.0982 | -0.47956
0.58601 | -34.0993 | -0.48981
0.59211 | -34.101 | -0.49636
0.59822 | -34.1023 | -0.50597
0.60432 | -34.1038 | -0.51818
0.61043 | -34.1051 | -0.52632
0.61653 | -34.1063 | -0.53674
0.62263 | -34.1078 | -0.54626
0.62874 | -34.1085 | -0.55702
0.63484 | -34.1093 | -0.56643
0.64095 | -34.1098 | -0.57674
0.64705 | -34.1102 | -0.58663
0.65316 | -34.1104 | -0.59914
0.65926 | -34.1105 | -0.61492
0.66536 | -34.1094 | -0.63219
0.67147 | -34.109 | -0.64538
0.67757 | -34.1086 | -0.65968
0.68368 | -34.1079 | -0.67845
0.68978 | -34.1071 | -0.69558
0.69589 | -34.1061 | -0.71442
0.70199 | -34.1051 | -0.73694
0.70809 | -34.1042 | -0.75595
0.7142 | -34.1035 | -0.77618
0.7203 | -34.1024 | -0.79525
0.72641 | -34.1017 | -0.81749
0.73251 | -34.101 | -0.8373
0.73862 | -34.1012 | -0.85888
0.74472 | -34.1009 | -0.87875
0.75082 | -34.1006 | -0.90068
0.75693 | -34.1009 | -0.92233
0.76303 | -34.1014 | -0.94411
0.76914 | -34.1031 | -0.96436
0.77524 | -34.1041 | -0.98382
0.78135 | -34.1053 | -1.0037
0.78745 | -34.1066 | -1.0235
0.79355 | -34.1078 | -1.0431
0.79966 | -34.1095 | -1.0624
0.80576 | -34.1117 | -1.0818
0.81187 | -34.1133 | -1.1031
0.81797 | -34.1152 | -1.1242
0.82408 | -34.1168 | -1.1445
0.83018 | -34.1185 | -1.1648
0.83628 | -34.1204 | -1.1831
0.84239 | -34.1215 | -1.2034
0.84849 | -34.1229 | -1.2249
0.8546 | -34.1243 | -1.2445
0.8607 | -34.1256 | -1.2651
0.86681 | -34.1266 | -1.2838
0.87291 | -34.1276 | -1.3025
0.87901 | -34.128 | -1.3204
0.88512 | -34.1287 | -1.3383
0.89122 | -34.1294 | -1.3565
0.89733 | -34.1298 | -1.372
0.90343 | -34.1302 | -1.3894
0.90953 | -34.1304 | -1.4037
0.91564 | -34.1304 | -1.4195
0.92174 | -34.1306 | -1.4344
0.92785 | -34.1307 | -1.4489
0.93395 | -34.1308 | -1.4627
0.94006 | -34.1309 | -1.4755
0.94616 | -34.1306 | -1.4873
0.95226 | -34.1307 | -1.4999
0.95837 | -34.1308 | -1.5136
0.96447 | -34.131 | -1.5263
0.97058 | -34.1308 | -1.5396
0.97668 | -34.1306 | -1.5528
0.98279 | -34.1303 | -1.5654
0.98889 | -34.1301 | -1.5759
0.99499 | -34.1301 | -1.5891
1.0011 | -34.13 | -1.5996
1.0072 | -34.1296 | -1.6111
1.0133 | -34.1293 | -1.6217
1.0194 | -34.1288 | -1.6344
1.0255 | -34.1285 | -1.6479
1.0316 | -34.1279 | -1.6596
1.0377 | -34.1275 | -1.673
1.0438 | -34.1269 | -1.6855
1.0499 | -34.1262 | -1.7005
1.056 | -34.1255 | -1.7145
1.0621 | -34.1248 | -1.7285
1.0682 | -34.1244 | -1.7407
1.0743 | -34.1237 | -1.7558
1.0805 | -34.1232 | -1.7699
1.0866 | -34.1223 | -1.7847
1.0927 | -34.1214 | -1.7989
1.0988 | -34.1208 | -1.8135
1.1049 | -34.12 | -1.8275
1.111 | -34.1191 | -1.8425
1.1171 | -34.1184 | -1.8568
1.1232 | -34.118 | -1.8716
1.1293 | -34.1175 | -1.8854
1.1354 | -34.1172 | -1.9001
1.1415 | -34.1168 | -1.9153
1.1476 | -34.1162 | -1.9305
1.1537 | -34.1157 | -1.9449
1.1598 | -34.1158 | -1.9593
1.1659 | -34.1155 | -1.9723
1.172 | -34.1152 | -1.986
1.1781 | -34.115 | -1.9996
1.1842 | -34.1148 | -2.0135
1.1903 | -34.1152 | -2.0243
1.1964 | -34.1151 | -2.0375
1.2025 | -34.1152 | -2.0511
1.2086 | -34.1155 | -2.0632
1.2147 | -34.1155 | -2.0755
1.2209 | -34.1158 | -2.0864
1.227 | -34.1165 | -2.097
1.2331 | -34.1169 | -2.1087
1.2392 | -34.1172 | -2.1203
1.2453 | -34.1177 | -2.1312
1.2514 | -34.1181 | -2.1409
1.2575 | -34.1186 | -2.1517
1.2636 | -34.1195 | -2.1605
1.2697 | -34.12 | -2.1696
1.2758 | -34.1207 | -2.1806
1.2819 | -34.1214 | -2.1903
1.288 | -34.122 | -2.2009
1.2941 | -34.1228 | -2.2092
1.3002 | -34.1235 | -2.2172
1.3063 | -34.124 | -2.226
1.3124 | -34.1247 | -2.2349
1.3185 | -34.1254 | -2.2439
1.3246 | -34.1261 | -2.2528
1.3307 | -34.1269 | -2.2613
1.3368 | -34.1277 | -2.2698
1.3429 | -34.1283 | -2.2786
1.349 | -34.1288 | -2.2878
1.3551 | -34.1295 | -2.297
1.3613 | -34.13 | -2.3052
1.3674 | -34.1306 | -2.3151
1.3735 | -34.131 | -2.3246
1.3796 | -34.1315 | -2.3336
1.3857 | -34.1318 | -2.3427
1.3918 | -34.1323 | -2.3527
1.3979 | -34.1329 | -2.3617
1.404 | -34.1333 | -2.3728
1.4101 | -34.1337 | -2.383
1.4162 | -34.1341 | -2.3934
1.4223 | -34.1346 | -2.4046
1.4284 | -34.135 | -2.4144
1.4345 | -34.1353 | -2.4257
1.4406 | -34.1355 | -2.4389
1.4467 | -34.136 | -2.4503
1.4528 | -34.1364 | -2.4614
1.4589 | -34.1369 | -2.4731
1.465 | -34.1371 | -2.4837
1.4711 | -34.1376 | -2.4948
1.4772 | -34.1381 | -2.5079
1.4833 | -34.1385 | -2.5206
1.4894 | -34.139 | -2.5326
1.4955 | -34.1395 | -2.5443
1.5016 | -34.14 | -2.5565
1.5078 | -34.1404 | -2.5701
1.5139 | -34.1408 | -2.5829
1.52 | -34.1412 | -2.5952
1.5261 | -34.1415 | -2.6076
1.5322 | -34.1418 | -2.6213
1.5383 | -34.1422 | -2.6336
1.5444 | -34.1427 | -2.646
1.5505 | -34.1432 | -2.6591
1.5566 | -34.1437 | -2.6718
1.5627 | -34.1441 | -2.6853
1.5688 | -34.1447 | -2.6986
1.5749 | -34.145 | -2.7109
1.581 | -34.1453 | -2.7239
1.5871 | -34.1457 | -2.7374
1.5932 | -34.146 | -2.75
1.5993 | -34.1464 | -2.7639
1.6054 | -34.1468 | -2.7768
1.6115 | -34.1472 | -2.7906
1.6176 | -34.1474 | -2.8029
1.6237 | -34.1476 | -2.8162
1.6298 | -34.1479 | -2.8285
1.6359 | -34.1483 | -2.8409
1.642 | -34.1486 | -2.8526
1.6482 | -34.1488 | -2.8652
1.6543 | -34.1492 | -2.8786
1.6604 | -34.1495 | -2.8903
1.6665 | -34.1496 | -2.9025
1.6726 | -34.1498 | -2.9147
1.6787 | -34.1501 | -2.927
1.6848 | -34.15 | -2.9392
1.6909 | -34.1502 | -2.9516
1.697 | -34.1504 | -2.9635
1.7031 | -34.1505 | -2.9758
1.7092 | -34.1504 | -2.9876
1.7153 | -34.1506 | -2.9993
1.7214 | -34.1503 | -3.0113
1.7275 | -34.1503 | -3.0236
1.7336 | -34.1502 | -3.0363
1.7397 | -34.1503 | -3.0483
1.7458 | -34.1503 | -3.0602
1.7519 | -34.1503 | -3.0721
1.758 | -34.1503 | -3.0838
1.7641 | -34.1505 | -3.0951
1.7702 | -34.1504 | -3.1078
1.7763 | -34.1503 | -3.119
1.7824 | -34.1503 | -3.1311
1.7885 | -34.1502 | -3.1434
1.7947 | -34.1503 | -3.1551
1.8008 | -34.1503 | -3.1661
1.8069 | -34.1503 | -3.1774
1.813 | -34.1504 | -3.1894
1.8191 | -34.1503 | -3.2018
1.8252 | -34.1504 | -3.2144
1.8313 | -34.1505 | -3.2264
1.8374 | -34.1506 | -3.238
1.8435 | -34.1506 | -3.2497
1.8496 | -34.1507 | -3.2625
1.8557 | -34.1508 | -3.275
1.8618 | -34.1511 | -3.2871
1.8679 | -34.1511 | -3.2989
1.874 | -34.1512 | -3.3107
1.8801 | -34.1514 | -3.3229
1.8862 | -34.1516 | -3.3346
1.8923 | -34.1517 | -3.3466
1.8984 | -34.1518 | -3.3585
1.9045 | -34.1518 | -3.3699
1.9106 | -34.1519 | -3.3819
1.9167 | -34.1519 | -3.394
1.9228 | -34.152 | -3.4056
1.9289 | -34.1521 | -3.4174
1.9351 | -34.1522 | -3.4281
1.9412 | -34.1524 | -3.4392
1.9473 | -34.1527 | -3.4511
1.9534 | -34.1529 | -3.4623
1.9595 | -34.153 | -3.474
1.9656 | -34.1533 | -3.4862
1.9717 | -34.1535 | -3.4977
1.9778 | -34.1538 | -3.509
1.9839 | -34.154 | -3.5203
1.99 | -34.1542 | -3.5312
1.9961 | -34.1542 | -3.5432
2.0022 | -34.1545 | -3.5545
2.0083 | -34.1545 | -3.5656
2.0144 | -34.1547 | -3.5769
2.0205 | -34.155 | -3.5885
2.0266 | -34.1553 | -3.5997
2.0327 | -34.1555 | -3.6112
2.0388 | -34.156 | -3.6226
2.0449 | -34.1561 | -3.6339
2.051 | -34.1564 | -3.6451
2.0571 | -34.1567 | -3.6563
2.0632 | -34.1569 | -3.6676
2.0693 | -34.157 | -3.6796
2.0754 | -34.1573 | -3.6908
2.0816 | -34.1576 | -3.7028
2.0877 | -34.1578 | -3.7141
2.0938 | -34.1581 | -3.7261
2.0999 | -34.1583 | -3.7376
2.106 | -34.1585 | -3.7491
2.1121 | -34.1587 | -3.7608
2.1182 | -34.1589 | -3.7715
2.1243 | -34.1592 | -3.782
2.1304 | -34.1593 | -3.793
2.1365 | -34.1596 | -3.8046
2.1426 | -34.1599 | -3.8164
2.1487 | -34.1601 | -3.8274
2.1548 | -34.1605 | -3.8391
2.1609 | -34.161 | -3.8503
2.167 | -34.1613 | -3.8613
2.1731 | -34.1616 | -3.8731
2.1792 | -34.1618 | -3.884
2.1853 | -34.1622 | -3.8958
2.1914 | -34.1625 | -3.9069
2.1975 | -34.1629 | -3.9183
2.2036 | -34.1631 | -3.9302
2.2097 | -34.1633 | -3.9415
2.2158 | -34.1637 | -3.9519
2.222 | -34.164 | -3.9622
2.2281 | -34.1642 | -3.9734
2.2342 | -34.1645 | -3.9846
2.2403 | -34.1647 | -3.995
2.2464 | -34.165 | -4.0056
2.2525 | -34.1652 | -4.0168
2.2586 | -34.1655 | -4.0272
2.2647 | -34.1658 | -4.0374
2.2708 | -34.1661 | -4.0484
2.2769 | -34.1665 | -4.0589
2.283 | -34.1667 | -4.0699
2.2891 | -34.1669 | -4.0806
2.2952 | -34.167 | -4.0909
2.3013 | -34.1673 | -4.1016
2.3074 | -34.1675 | -4.1119
2.3135 | -34.1677 | -4.123
2.3196 | -34.168 | -4.133
2.3257 | -34.1682 | -4.1435
2.3318 | -34.1683 | -4.1545
2.3379 | -34.1685 | -4.165
2.344 | -34.1686 | -4.1762
2.3501 | -34.1688 | -4.1868
2.3562 | -34.169 | -4.1969
2.3623 | -34.1691 | -4.2072
2.3685 | -34.1692 | -4.2175
2.3746 | -34.1694 | -4.2285
2.3807 | -34.1696 | -4.2388
2.3868 | -34.1697 | -4.2493
2.3929 | -34.1699 | -4.2598
2.399 | -34.17 | -4.2701
2.4051 | -34.1703 | -4.2804
2.4112 | -34.1704 | -4.2909
2.4173 | -34.1706 | -4.3013
2.4234 | -34.1707 | -4.3119
2.4295 | -34.1708 | -4.3226
2.4356 | -34.171 | -4.3332
2.4417 | -34.1711 | -4.3442
2.4478 | -34.1712 | -4.3548
2.4539 | -34.1714 | -4.3649
2.46 | -34.1715 | -4.3761
2.4661 | -34.1716 | -4.387
2.4722 | -34.1718 | -4.397
2.4783 | -34.1719 | -4.4072
2.4844 | -34.1719 | -4.4181
2.4905 | -34.172 | -4.4289
2.4966 | -34.1722 | -4.4395
2.5027 | -34.1723 | -4.4497
2.5089 | -34.1724 | -4.4604
2.515 | -34.1724 | -4.471
2.5211 | -34.1725 | -4.4816
2.5272 | -34.1725 | -4.4923
2.5333 | -34.1727 | -4.503
2.5394 | -34.1728 | -4.5136
2.5455 | -34.1729 | -4.5248
2.5516 | -34.1731 | -4.5349
2.5577 | -34.1732 | -4.5462
2.5638 | -34.1734 | -4.5576
2.5699 | -34.1734 | -4.5688
2.576 | -34.1735 | -4.5802
2.5821 | -34.1736 | -4.5912
2.5882 | -34.1737 | -4.6021
2.5943 | -34.1738 | -4.6128
2.6004 | -34.1739 | -4.6235
2.6065 | -34.174 | -4.6337
2.6126 | -34.1741 | -4.6439
2.6187 | -34.1743 | -4.6542
2.6248 | -34.1744 | -4.6647
2.6309 | -34.1745 | -4.6759
2.637 | -34.1748 | -4.6873
2.6431 | -34.1749 | -4.6976
2.6492 | -34.175 | -4.7087
2.6554 | -34.1752 | -4.7193
2.6615 | -34.1752 | -4.7301
2.6676 | -34.1755 | -4.7412
2.6737 | -34.1757 | -4.7522
2.6798 | -34.176 | -4.763
2.6859 | -34.1762 | -4.774
2.692 | -34.1763 | -4.7847
2.6981 | -34.1765 | -4.7956
2.7042 | -34.1767 | -4.8065
2.7103 | -34.1768 | -4.8169
2.7164 | -34.1771 | -4.8277
2.7225 | -34.1773 | -4.8388
2.7286 | -34.1775 | -4.8495
2.7347 | -34.1777 | -4.8598
2.7408 | -34.178 | -4.8697
2.7469 | -34.1782 | -4.8808
2.753 | -34.1784 | -4.8912
2.7591 | -34.1787 | -4.9022
2.7652 | -34.1789 | -4.9134
2.7713 | -34.1791 | -4.9239
2.7774 | -34.1793 | -4.9346
2.7835 | -34.1795 | -4.9454
2.7896 | -34.1797 | -4.9557
2.7958 | -34.1799 | -4.9662
2.8019 | -34.1802 | -4.9764
2.808 | -34.1804 | -4.9869
2.8141 | -34.1807 | -4.9978
2.8202 | -34.181 | -5.0085
2.8263 | -34.1813 | -5.0191
2.8324 | -34.1816 | -5.0302
2.8385 | -34.182 | -5.0409
2.8446 | -34.1823 | -5.0514
2.8507 | -34.1827 | -5.062
2.8568 | -34.183 | -5.0722
2.8629 | -34.1833 | -5.0826
2.869 | -34.1836 | -5.0935
2.8751 | -34.1839 | -5.104
2.8812 | -34.1843 | -5.1151
2.8873 | -34.1846 | -5.1257
2.8934 | -34.1849 | -5.1365
2.8995 | -34.1852 | -5.1474
2.9056 | -34.1854 | -5.1583
2.9117 | -34.1857 | -5.1691
2.9178 | -34.1861 | -5.1796
2.9239 | -34.1865 | -5.1909
2.93 | -34.1869 | -5.2022
2.9361 | -34.1872 | -5.2128
2.9423 | -34.1875 | -5.2232
2.9484 | -34.1879 | -5.2341
2.9545 | -34.1882 | -5.2449
2.9606 | -34.1886 | -5.256
2.9667 | -34.189 | -5.267
2.9728 | -34.1893 | -5.278
2.9789 | -34.1897 | -5.2889
2.985 | -34.1901 | -5.3
2.9911 | -34.1905 | -5.3111
2.9972 | -34.1909 | -5.322
3.0033 | -34.1913 | -5.3333
3.0094 | -34.1917 | -5.3441
3.0155 | -34.1921 | -5.3549
3.0216 | -34.1924 | -5.3659
3.0277 | -34.1928 | -5.3774
3.0338 | -34.1931 | -5.3886
3.0399 | -34.1935 | -5.3997
3.046 | -34.1938 | -5.4111
3.0521 | -34.1942 | -5.422
3.0582 | -34.1944 | -5.4339
3.0643 | -34.1948 | -5.4452
3.0704 | -34.1951 | -5.4568
3.0765 | -34.1954 | -5.4677
3.0827 | -34.1956 | -5.4787
3.0888 | -34.196 | -5.4898
3.0949 | -34.1964 | -5.501
3.101 | -34.1967 | -5.5123
3.1071 | -34.197 | -5.5238
3.1132 | -34.1974 | -5.5354
3.1193 | -34.1977 | -5.5467
3.1254 | -34.1981 | -5.5579
3.1315 | -34.1984 | -5.5696
3.1376 | -34.1987 | -5.5808
3.1437 | -34.199 | -5.5922
3.1498 | -34.1994 | -5.604
3.1559 | -34.1997 | -5.6153
3.162 | -34.1999 | -5.6272
3.1681 | -34.2002 | -5.639
3.1742 | -34.2005 | -5.6504
3.1803 | -34.2008 | -5.6619
3.1864 | -34.2012 | -5.673
3.1925 | -34.2015 | -5.6844
3.1986 | -34.2017 | -5.6962
3.2047 | -34.202 | -5.7081
3.2108 | -34.2022 | -5.7191
3.2169 | -34.2025 | -5.7306
3.223 | -34.2028 | -5.7422
3.2292 | -34.203 | -5.7538
3.2353 | -34.2032 | -5.7645
3.2414 | -34.2034 | -5.7758
3.2475 | -34.2036 | -5.787
3.2536 | -34.2039 | -5.7983
3.2597 | -34.2041 | -5.8091
3.2658 | -34.2043 | -5.8199
3.2719 | -34.2045 | -5.8308
3.278 | -34.2048 | -5.8423
3.2841 | -34.205 | -5.8532
3.2902 | -34.2053 | -5.8645
3.2963 | -34.2055 | -5.8758
3.3024 | -34.2056 | -5.8868
3.3085 | -34.2059 | -5.8982
3.3146 | -34.206 | -5.9092
3.3207 | -34.2063 | -5.9205
3.3268 | -34.2065 | -5.9318
3.3329 | -34.2067 | -5.9424
3.339 | -34.2068 | -5.9531
3.3451 | -34.207 | -5.9642
3.3512 | -34.2072 | -5.9747
3.3573 | -34.2074 | -5.9858
3.3634 | -34.2076 | -5.9968
3.3696 | -34.2078 | -6.008
3.3757 | -34.2079 | -6.0187
3.3818 | -34.2081 | -6.0294
3.3879 | -34.2083 | -6.0404
3.394 | -34.2084 | -6.0516
3.4001 | -34.2086 | -6.0623
3.4062 | -34.2087 | -6.0731
3.4123 | -34.2088 | -6.0835
3.4184 | -34.209 | -6.0946
3.4245 | -34.2092 | -6.1053
3.4306 | -34.2093 | -6.1158
3.4367 | -34.2095 | -6.1262
3.4428 | -34.2096 | -6.1365
3.4489 | -34.2097 | -6.1473
3.455 | -34.2099 | -6.158
3.4611 | -34.2101 | -6.1691
3.4672 | -34.2102 | -6.18
3.4733 | -34.2104 | -6.1909
3.4794 | -34.2105 | -6.2015
3.4855 | -34.2107 | -6.2118
3.4916 | -34.2108 | -6.222
3.4977 | -34.211 | -6.2331
3.5038 | -34.2111 | -6.2439
3.5099 | -34.2112 | -6.254
3.5161 | -34.2114 | -6.2642
3.5222 | -34.2115 | -6.2744
3.5283 | -34.2116 | -6.2851
3.5344 | -34.2117 | -6.2957
3.5405 | -34.2118 | -6.3061
3.5466 | -34.212 | -6.3169
3.5527 | -34.2121 | -6.3275
3.5588 | -34.2122 | -6.3379
3.5649 | -34.2123 | -6.3484
3.571 | -34.2124 | -6.3589
3.5771 | -34.2125 | -6.3698
3.5832 | -34.2126 | -6.3796
3.5893 | -34.2128 | -6.39
3.5954 | -34.213 | -6.4007
3.6015 | -34.2132 | -6.4114
3.6076 | -34.2134 | -6.4218
3.6137 | -34.2135 | -6.4317
3.6198 | -34.2137 | -6.4418
3.6259 | -34.2139 | -6.4525
3.632 | -34.214 | -6.463
3.6381 | -34.2142 | -6.4736
3.6442 | -34.2143 | -6.484
3.6503 | -34.2145 | -6.494
3.6565 | -34.2147 | -6.5046
3.6626 | -34.2148 | -6.515
3.6687 | -34.215 | -6.5256
3.6748 | -34.2153 | -6.5358
3.6809 | -34.2155 | -6.5463
3.687 | -34.2157 | -6.5571
3.6931 | -34.2159 | -6.5672
3.6992 | -34.2161 | -6.5776
3.7053 | -34.2163 | -6.5879
3.7114 | -34.2164 | -6.5985
3.7175 | -34.2167 | -6.609
3.7236 | -34.2169 | -6.6189
3.7297 | -34.2171 | -6.6295
3.7358 | -34.2173 | -6.6402
3.7419 | -34.2175 | -6.6509
3.748 | -34.2177 | -6.6613
3.7541 | -34.2179 | -6.6717
3.7602 | -34.2181 | -6.682
3.7663 | -34.2183 | -6.6923
3.7724 | -34.2186 | -6.7028
3.7785 | -34.2189 | -6.713
3.7846 | -34.219 | -6.7233
3.7907 | -34.2193 | -6.734
3.7969 | -34.2196 | -6.7446
3.803 | -34.2198 | -6.7546
3.8091 | -34.22 | -6.765
3.8152 | -34.2203 | -6.7753
3.8213 | -34.2205 | -6.7859
3.8274 | -34.2208 | -6.7961
3.8335 | -34.2211 | -6.8066
3.8396 | -34.2214 | -6.8171
3.8457 | -34.2217 | -6.8276
3.8518 | -34.222 | -6.838
3.8579 | -34.2223 | -6.8484
3.864 | -34.2226 | -6.859
3.8701 | -34.2229 | -6.8698
3.8762 | -34.2231 | -6.8801
3.8823 | -34.2234 | -6.8907
3.8884 | -34.2237 | -6.9012
3.8945 | -34.224 | -6.9119
3.9006 | -34.2243 | -6.9227
3.9067 | -34.2246 | -6.9336
3.9128 | -34.2249 | -6.9442
3.9189 | -34.2252 | -6.9544
3.925 | -34.2254 | -6.965
3.9311 | -34.2257 | -6.9751
3.9372 | -34.226 | -6.9858
3.9434 | -34.2262 | -6.9963
3.9495 | -34.2266 | -7.007
3.9556 | -34.2269 | -7.0173
3.9617 | -34.2272 | -7.0274
3.9678 | -34.2275 | -7.0376
3.9739 | -34.2278 | -7.0481
3.98 | -34.2281 | -7.0587
3.9861 | -34.2283 | -7.0691
3.9922 | -34.2286 | -7.0796
3.9983 | -34.2289 | -7.0897
4.0044 | -34.2292 | -7.0998
4.0105 | -34.2294 | -7.1103
4.0166 | -34.2298 | -7.1208
4.0227 | -34.2301 | -7.1313
4.0288 | -34.2303 | -7.1422
4.0349 | -34.2305 | -7.1524
4.041 | -34.2308 | -7.1629
4.0471 | -34.231 | -7.1734
4.0532 | -34.2313 | -7.1835
4.0593 | -34.2316 | -7.1942
4.0654 | -34.2319 | -7.2044
4.0715 | -34.2322 | -7.2153
4.0776 | -34.2325 | -7.2257
4.0838 | -34.2328 | -7.2366
4.0899 | -34.2331 | -7.2468
4.096 | -34.2333 | -7.2571
4.1021 | -34.2336 | -7.2674
4.1082 | -34.2338 | -7.2776
4.1143 | -34.2341 | -7.2878
4.1204 | -34.2344 | -7.2979
4.1265 | -34.2347 | -7.308
4.1326 | -34.2349 | -7.3181
4.1387 | -34.2353 | -7.3283
4.1448 | -34.2355 | -7.3388
4.1509 | -34.2357 | -7.3492
4.157 | -34.236 | -7.36
4.1631 | -34.2362 | -7.3706
4.1692 | -34.2365 | -7.3808
4.1753 | -34.2368 | -7.3911
4.1814 | -34.237 | -7.4011
4.1875 | -34.2373 | -7.4118
4.1936 | -34.2375 | -7.4223
4.1997 | -34.2378 | -7.4323
4.2058 | -34.238 | -7.4429
4.2119 | -34.2383 | -7.4531
4.218 | -34.2385 | -7.4632
4.2241 | -34.2387 | -7.4733
4.2303 | -34.239 | -7.4835
4.2364 | -34.2392 | -7.4938
4.2425 | -34.2394 | -7.5041
4.2486 | -34.2396 | -7.5146
4.2547 | -34.2398 | -7.5246
4.2608 | -34.24 | -7.535
4.2669 | -34.2402 | -7.545
4.273 | -34.2404 | -7.5554
4.2791 | -34.2406 | -7.5657
4.2852 | -34.2409 | -7.5762
4.2913 | -34.241 | -7.5861
4.2974 | -34.2412 | -7.5967
4.3035 | -34.2415 | -7.6072
4.3096 | -34.2417 | -7.6177
4.3157 | -34.2419 | -7.6278
4.3218 | -34.2421 | -7.6379
4.3279 | -34.2423 | -7.6482
4.334 | -34.2425 | -7.6584
4.3401 | -34.2427 | -7.6682
4.3462 | -34.243 | -7.6786
4.3523 | -34.2432 | -7.6884
4.3584 | -34.2434 | -7.6982
4.3645 | -34.2436 | -7.7083
4.3707 | -34.2438 | -7.7182
4.3768 | -34.244 | -7.7284
4.3829 | -34.2442 | -7.739
4.389 | -34.2444 | -7.7491
4.3951 | -34.2446 | -7.7591
4.4012 | -34.2448 | -7.7696
4.4073 | -34.245 | -7.7796
4.4134 | -34.2452 | -7.7897
4.4195 | -34.2455 | -7.7998
4.4256 | -34.2457 | -7.8097
4.4317 | -34.2459 | -7.8198
4.4378 | -34.2461 | -7.8296
4.4439 | -34.2463 | -7.8395
4.45 | -34.2464 | -7.8496
4.4561 | -34.2466 | -7.8597
4.4622 | -34.2468 | -7.8699
4.4683 | -34.247 | -7.8799
4.4744 | -34.2472 | -7.8904
4.4805 | -34.2474 | -7.9001
4.4866 | -34.2476 | -7.9102
4.4927 | -34.2478 | -7.9205
4.4988 | -34.248 | -7.9305
4.5049 | -34.2482 | -7.9402
4.511 | -34.2484 | -7.95
4.5172 | -34.2486 | -7.9601
4.5233 | -34.2488 | -7.9705
4.5294 | -34.249 | -7.9804
4.5355 | -34.2492 | -7.9905
4.5416 | -34.2494 | -8.0005
4.5477 | -34.2496 | -8.0106
4.5538 | -34.2498 | -8.0202
4.5599 | -34.25 | -8.0302
4.566 | -34.2502 | -8.0402
4.5721 | -34.2504 | -8.0503
4.5782 | -34.2507 | -8.0606
4.5843 | -34.2509 | -8.0707
4.5904 | -34.2511 | -8.0808
4.5965 | -34.2513 | -8.0908
4.6026 | -34.2515 | -8.1013
4.6087 | -34.2517 | -8.1112
4.6148 | -34.2519 | -8.1216
4.6209 | -34.2521 | -8.1316
4.627 | -34.2524 | -8.1418
4.6331 | -34.2525 | -8.1518
4.6392 | -34.2527 | -8.1619
4.6453 | -34.2529 | -8.1719
4.6514 | -34.2532 | -8.182
4.6576 | -34.2533 | -8.1925
4.6637 | -34.2535 | -8.2022
4.6698 | -34.2537 | -8.2121
4.6759 | -34.2539 | -8.2225
4.682 | -34.2541 | -8.2324
4.6881 | -34.2542 | -8.2425
4.6942 | -34.2544 | -8.2525
4.7003 | -34.2546 | -8.2626
4.7064 | -34.2548 | -8.2725
4.7125 | -34.2551 | -8.2828
4.7186 | -34.2553 | -8.2926
4.7247 | -34.2555 | -8.3027
4.7308 | -34.2557 | -8.3127
4.7369 | -34.2559 | -8.3229
4.743 | -34.2561 | -8.3333
4.7491 | -34.2564 | -8.3437
4.7552 | -34.2565 | -8.3539
4.7613 | -34.2567 | -8.364
4.7674 | -34.2569 | -8.374
4.7735 | -34.2571 | -8.3845
4.7796 | -34.2574 | -8.3949
4.7857 | -34.2576 | -8.4051
4.7918 | -34.2578 | -8.415
4.7979 | -34.258 | -8.425
4.8041 | -34.2582 | -8.4351
4.8102 | -34.2585 | -8.4449
4.8163 | -34.2587 | -8.4549
4.8224 | -34.2589 | -8.4651
4.8285 | -34.2591 | -8.4749
4.8346 | -34.2594 | -8.4848
4.8407 | -34.2595 | -8.4949
4.8468 | -34.2598 | -8.5053
4.8529 | -34.26 | -8.5157
4.859 | -34.2602 | -8.5261
4.8651 | -34.2605 | -8.536
4.8712 | -34.2607 | -8.5459
4.8773 | -34.2609 | -8.556
4.8834 | -34.2611 | -8.5662
4.8895 | -34.2613 | -8.5761
4.8956 | -34.2615 | -8.5862
4.9017 | -34.2617 | -8.5961
4.9078 | -34.2619 | -8.6065
4.9139 | -34.2621 | -8.6162
4.92 | -34.2624 | -8.626
4.9261 | -34.2627 | -8.6358
4.9322 | -34.2629 | -8.6459
4.9383 | -34.2631 | -8.656
4.9445 | -34.2634 | -8.6661
4.9506 | -34.2636 | -8.676
4.9567 | -34.2638 | -8.6859
4.9628 | -34.264 | -8.6959
4.9689 | -34.2643 | -8.7064
4.975 | -34.2645 | -8.7168
4.9811 | -34.2647 | -8.7266
4.9872 | -34.265 | -8.7366
4.9933 | -34.2653 | -8.7469
4.9994 | -34.2655 | -8.7574
5.0055 | -34.2658 | -8.7675
5.0116 | -34.266 | -8.7781
5.0177 | -34.2663 | -8.7884
5.0238 | -34.2666 | -8.7988
5.0299 | -34.2669 | -8.8093
5.036 | -34.2671 | -8.8194
5.0421 | -34.2673 | -8.8295
5.0482 | -34.2676 | -8.8396
5.0543 | -34.2678 | -8.8495
5.0604 | -34.2681 | -8.8596
5.0665 | -34.2684 | -8.8695
5.0726 | -34.2686 | -8.8799
5.0787 | -34.2689 | -8.8903
5.0848 | -34.2692 | -8.9006
5.091 | -34.2694 | -8.9105
5.0971 | -34.2697 | -8.9203
5.1032 | -34.27 | -8.9308
5.1093 | -34.2702 | -8.941
5.1154 | -34.2704 | -8.9509
5.1215 | -34.2707 | -8.9608
5.1276 | -34.2709 | -8.9707
5.1337 | -34.2713 | -8.9806
5.1398 | -34.2715 | -8.9905
5.1459 | -34.2718 | -9.0009
5.152 | -34.272 | -9.0112
5.1581 | -34.2723 | -9.0215
5.1642 | -34.2725 | -9.0317
5.1703 | -34.2728 | -9.0419
5.1764 | -34.2731 | -9.0519
5.1825 | -34.2734 | -9.0618
5.1886 | -34.2736 | -9.0716
5.1947 | -34.2739 | -9.0818
5.2008 | -34.2742 | -9.0921
5.2069 | -34.2744 | -9.1025
5.213 | -34.2747 | -9.1124
5.2191 | -34.275 | -9.1226
5.2252 | -34.2752 | -9.1328
5.2314 | -34.2755 | -9.1431
5.2375 | -34.2758 | -9.1531
5.2436 | -34.276 | -9.1633
5.2497 | -34.2763 | -9.1732
5.2558 | -34.2766 | -9.1837
5.2619 | -34.2768 | -9.1941
5.268 | -34.2771 | -9.2045
5.2741 | -34.2775 | -9.2147
5.2802 | -34.2777 | -9.2253
5.2863 | -34.278 | -9.2354
5.2924 | -34.2783 | -9.2457
5.2985 | -34.2786 | -9.256
5.3046 | -34.2789 | -9.2663
5.3107 | -34.2792 | -9.2767
5.3168 | -34.2795 | -9.287
5.3229 | -34.2797 | -9.2972
5.329 | -34.28 | -9.3074
5.3351 | -34.2803 | -9.3173
5.3412 | -34.2805 | -9.3275
5.3473 | -34.2808 | -9.3378
5.3534 | -34.2812 | -9.3482
5.3595 | -34.2814 | -9.3584
5.3656 | -34.2817 | -9.3687
5.3717 | -34.282 | -9.3792
5.3779 | -34.2823 | -9.3898
5.384 | -34.2826 | -9.4
5.3901 | -34.2829 | -9.4105
5.3962 | -34.2831 | -9.4206
5.4023 | -34.2834 | -9.4311
5.4084 | -34.2838 | -9.4413
5.4145 | -34.2841 | -9.4513
5.4206 | -34.2843 | -9.4616
5.4267 | -34.2846 | -9.4722
5.4328 | -34.2849 | -9.4825
5.4389 | -34.2852 | -9.4925
5.445 | -34.2855 | -9.5031
5.4511 | -34.2858 | -9.5132
5.4572 | -34.2861 | -9.5235
5.4633 | -34.2864 | -9.5338
5.4694 | -34.2866 | -9.5441
5.4755 | -34.2869 | -9.5541
5.4816 | -34.2872 | -9.5644
5.4877 | -34.2875 | -9.5746
5.4938 | -34.2878 | -9.5846
5.4999 | -34.2881 | -9.5948
5.506 | -34.2884 | -9.605
5.5121 | -34.2887 | -9.6152
5.5183 | -34.289 | -9.6253
5.5244 | -34.2893 | -9.6358
5.5305 | -34.2896 | -9.646
5.5366 | -34.2899 | -9.656
5.5427 | -34.2901 | -9.6664
5.5488 | -34.2904 | -9.6768
5.5549 | -34.2907 | -9.6871
5.561 | -34.291 | -9.6976
5.5671 | -34.2913 | -9.7077
5.5732 | -34.2916 | -9.7183
5.5793 | -34.2919 | -9.7289
5.5854 | -34.2922 | -9.7391
5.5915 | -34.2925 | -9.7497
5.5976 | -34.2928 | -9.7601
5.6037 | -34.2931 | -9.7704
5.6098 | -34.2933 | -9.7807
5.6159 | -34.2936 | -9.7908
5.622 | -34.2939 | -9.8012
5.6281 | -34.2942 | -9.8116
5.6342 | -34.2945 | -9.8222
5.6403 | -34.2948 | -9.8326
5.6464 | -34.2951 | -9.843
5.6525 | -34.2954 | -9.8534
5.6586 | -34.2957 | -9.8635
5.6648 | -34.296 | -9.8738
5.6709 | -34.2964 | -9.8843
5.677 | -34.2967 | -9.8946
5.6831 | -34.297 | -9.905
5.6892 | -34.2972 | -9.9155
5.6953 | -34.2975 | -9.9259
5.7014 | -34.2978 | -9.9363
5.7075 | -34.2981 | -9.9466
5.7136 | -34.2984 | -9.9569
5.7197 | -34.2987 | -9.9672
5.7258 | -34.2989 | -9.9773
5.7319 | -34.2993 | -9.9877
5.738 | -34.2996 | -9.998
5.7441 | -34.2999 | -10.0083
5.7502 | -34.3002 | -10.0183
5.7563 | -34.3005 | -10.0287
5.7624 | -34.3008 | -10.0392
5.7685 | -34.3011 | -10.0496
5.7746 | -34.3014 | -10.0602
5.7807 | -34.3017 | -10.0705
5.7868 | -34.3019 | -10.0811
5.7929 | -34.3022 | -10.0913
5.799 | -34.3025 | -10.1015
5.8052 | -34.3028 | -10.1117
5.8113 | -34.3031 | -10.1219
5.8174 | -34.3034 | -10.1321
5.8235 | -34.3037 | -10.1424
5.8296 | -34.304 | -10.1529
5.8357 | -34.3042 | -10.1631
5.8418 | -34.3045 | -10.1736
5.8479 | -34.3048 | -10.184
5.854 | -34.3051 | -10.1944
5.8601 | -34.3054 | -10.2046
5.8662 | -34.3057 | -10.2151
5.8723 | -34.3059 | -10.2252
5.8784 | -34.3062 | -10.2355
5.8845 | -34.3065 | -10.2456
5.8906 | -34.3068 | -10.256
5.8967 | -34.3071 | -10.2661
5.9028 | -34.3074 | -10.2763
5.9089 | -34.3078 | -10.2864
5.915 | -34.308 | -10.2965
5.9211 | -34.3083 | -10.3068
5.9272 | -34.3086 | -10.3168
5.9333 | -34.3089 | -10.327
5.9394 | -34.3092 | -10.3373
5.9455 | -34.3095 | -10.3473
5.9517 | -34.3098 | -10.3574
5.9578 | -34.3102 | -10.3678
5.9639 | -34.3105 | -10.3777
5.97 | -34.3108 | -10.388
5.9761 | -34.3111 | -10.3982
5.9822 | -34.3114 | -10.4085
5.9883 | -34.3117 | -10.4188
5.9944 | -34.312 | -10.4291
6.0005 | -34.3123 | -10.4395
6.0066 | -34.3126 | -10.4496
6.0127 | -34.3129 | -10.4595
6.0188 | -34.3132 | -10.4701
6.0249 | -34.3135 | -10.4804
6.031 | -34.3138 | -10.4904
6.0371 | -34.3141 | -10.5007
6.0432 | -34.3144 | -10.5108
6.0493 | -34.3147 | -10.5208
6.0554 | -34.315 | -10.5309
6.0615 | -34.3153 | -10.5411
6.0676 | -34.3157 | -10.5516
6.0737 | -34.316 | -10.5619
6.0798 | -34.3162 | -10.572
6.0859 | -34.3166 | -10.5822
6.0921 | -34.3169 | -10.5922
6.0982 | -34.3172 | -10.6024
6.1043 | -34.3175 | -10.6124
6.1104 | -34.3178 | -10.6228
6.1165 | -34.3182 | -10.6326
6.1226 | -34.3185 | -10.6427
6.1287 | -34.3188 | -10.6528
6.1348 | -34.3191 | -10.663
6.1409 | -34.3194 | -10.6731
6.147 | -34.3197 | -10.6832
6.1531 | -34.32 | -10.6933
6.1592 | -34.3203 | -10.7033
6.1653 | -34.3206 | -10.7133
6.1714 | -34.3209 | -10.7235
6.1775 | -34.3212 | -10.7338
6.1836 | -34.3215 | -10.7439
6.1897 | -34.3218 | -10.7541
6.1958 | -34.3221 | -10.7644
6.2019 | -34.3224 | -10.7747
6.208 | -34.3228 | -10.7849
6.2141 | -34.3231 | -10.7949
6.2202 | -34.3234 | -10.8052
6.2263 | -34.3237 | -10.8152
6.2325 | -34.324 | -10.825
6.2386 | -34.3243 | -10.8348
6.2447 | -34.3246 | -10.8452
6.2508 | -34.3249 | -10.8553
6.2569 | -34.3253 | -10.8653
6.263 | -34.3256 | -10.8754
6.2691 | -34.3259 | -10.8855
6.2752 | -34.3262 | -10.8956
6.2813 | -34.3264 | -10.9057
6.2874 | -34.3267 | -10.9157
6.2935 | -34.327 | -10.9257
6.2996 | -34.3273 | -10.9358
6.3057 | -34.3276 | -10.9459
6.3118 | -34.3279 | -10.9561
6.3179 | -34.3282 | -10.9661
6.324 | -34.3285 | -10.9762
6.3301 | -34.3288 | -10.9862
6.3362 | -34.3292 | -10.9962
6.3423 | -34.3295 | -11.0064
6.3484 | -34.3298 | -11.0165
6.3545 | -34.3301 | -11.0263
6.3606 | -34.3304 | -11.036
6.3667 | -34.3306 | -11.0459
6.3728 | -34.331 | -11.056
6.379 | -34.3313 | -11.0661
6.3851 | -34.3316 | -11.0759
6.3912 | -34.3318 | -11.0859
6.3973 | -34.3321 | -11.0959
6.4034 | -34.3324 | -11.1058
6.4095 | -34.3327 | -11.1158
6.4156 | -34.333 | -11.1258
6.4217 | -34.3333 | -11.1358
6.4278 | -34.3336 | -11.1459
6.4339 | -34.3339 | -11.1558
6.44 | -34.3343 | -11.1657
6.4461 | -34.3346 | -11.176
6.4522 | -34.3349 | -11.1857
6.4583 | -34.3351 | -11.1959
6.4644 | -34.3354 | -11.206
6.4705 | -34.3357 | -11.2156
6.4766 | -34.336 | -11.2257
6.4827 | -34.3363 | -11.2357
6.4888 | -34.3366 | -11.2455
6.4949 | -34.3369 | -11.2552
6.501 | -34.3372 | -11.2653
6.5071 | -34.3375 | -11.275
6.5132 | -34.3378 | -11.285
6.5194 | -34.3381 | -11.295
6.5255 | -34.3384 | -11.3048
6.5316 | -34.3387 | -11.3149
6.5377 | -34.339 | -11.3248
6.5438 | -34.3392 | -11.3345
6.5499 | -34.3396 | -11.3444
6.556 | -34.3399 | -11.3543
6.5621 | -34.3401 | -11.3645
6.5682 | -34.3404 | -11.3745
6.5743 | -34.3408 | -11.3843
6.5804 | -34.3411 | -11.3943
6.5865 | -34.3414 | -11.4043
6.5926 | -34.3417 | -11.4141
6.5987 | -34.342 | -11.4238
6.6048 | -34.3423 | -11.4337
6.6109 | -34.3426 | -11.4436
6.617 | -34.3429 | -11.4536
6.6231 | -34.3432 | -11.4635
6.6292 | -34.3435 | -11.4734
6.6353 | -34.3438 | -11.4836
6.6414 | -34.3441 | -11.4935
6.6475 | -34.3444 | -11.5033
6.6536 | -34.3447 | -11.5131
6.6597 | -34.345 | -11.5227
6.6659 | -34.3453 | -11.5327
6.672 | -34.3456 | -11.5424
6.6781 | -34.3459 | -11.5523
6.6842 | -34.3462 | -11.562
6.6903 | -34.3465 | -11.5719
6.6964 | -34.3468 | -11.5818
6.7025 | -34.3471 | -11.5915
6.7086 | -34.3474 | -11.6014
6.7147 | -34.3477 | -11.6113
6.7208 | -34.348 | -11.6213
6.7269 | -34.3483 | -11.6311
6.733 | -34.3486 | -11.6408
6.7391 | -34.3489 | -11.6506
6.7452 | -34.3492 | -11.6602
6.7513 | -34.3495 | -11.6698
6.7574 | -34.3498 | -11.6798
6.7635 | -34.3501 | -11.6894
6.7696 | -34.3504 | -11.6991
6.7757 | -34.3508 | -11.709
6.7818 | -34.3511 | -11.7188
6.7879 | -34.3514 | -11.7285
6.794 | -34.3517 | -11.7381
6.8001 | -34.3519 | -11.7479
6.8063 | -34.3522 | -11.7576
6.8124 | -34.3525 | -11.7676
6.8185 | -34.3528 | -11.7776
6.8246 | -34.3531 | -11.7872
6.8307 | -34.3534 | -11.7969
6.8368 | -34.3537 | -11.8069
6.8429 | -34.354 | -11.8167
6.849 | -34.3544 | -11.8266
6.8551 | -34.3547 | -11.8363
6.8612 | -34.355 | -11.846
6.8673 | -34.3553 | -11.8557
6.8734 | -34.3556 | -11.8656
6.8795 | -34.3559 | -11.8755
6.8856 | -34.3562 | -11.8852
6.8917 | -34.3565 | -11.895
6.8978 | -34.3568 | -11.9045
6.9039 | -34.3571 | -11.914
6.91 | -34.3574 | -11.9238
6.9161 | -34.3577 | -11.9333
6.9222 | -34.358 | -11.9431
6.9283 | -34.3584 | -11.9526
6.9344 | -34.3587 | -11.9624
6.9405 | -34.359 | -11.9721
6.9466 | -34.3593 | -11.982
6.9528 | -34.3597 | -11.9919
6.9589 | -34.36 | -12.0018
6.965 | -34.3603 | -12.0117
6.9711 | -34.3606 | -12.0215
6.9772 | -34.361 | -12.0312
6.9833 | -34.3613 | -12.041
6.9894 | -34.3616 | -12.0509
6.9955 | -34.3619 | -12.0604
7.0016 | -34.3622 | -12.0701
7.0077 | -34.3625 | -12.0799
7.0138 | -34.3629 | -12.0895
7.0199 | -34.3632 | -12.0992
7.026 | -34.3635 | -12.1088
7.0321 | -34.3638 | -12.1184
7.0382 | -34.3641 | -12.1278
7.0443 | -34.3645 | -12.1377
7.0504 | -34.3648 | -12.1475
7.0565 | -34.3651 | -12.1572
7.0626 | -34.3654 | -12.1667
7.0687 | -34.3657 | -12.176
7.0748 | -34.3661 | -12.1855
7.0809 | -34.3664 | -12.1951
7.087 | -34.3667 | -12.2049
7.0932 | -34.367 | -12.2146
7.0993 | -34.3673 | -12.2243
7.1054 | -34.3677 | -12.2338
7.1115 | -34.368 | -12.2433
7.1176 | -34.3683 | -12.2529
7.1237 | -34.3686 | -12.2625
7.1298 | -34.3689 | -12.2721
7.1359 | -34.3692 | -12.2818
7.142 | -34.3695 | -12.2912
7.1481 | -34.3699 | -12.3007
7.1542 | -34.3702 | -12.3104
7.1603 | -34.3705 | -12.3201
7.1664 | -34.3708 | -12.3297
7.1725 | -34.3712 | -12.3394
7.1786 | -34.3715 | -12.349
7.1847 | -34.3718 | -12.3587
7.1908 | -34.3721 | -12.3685
7.1969 | -34.3724 | -12.3783
7.203 | -34.3728 | -12.3878
7.2091 | -34.3731 | -12.3974
7.2152 | -34.3734 | -12.407
7.2213 | -34.3738 | -12.4165
7.2274 | -34.3741 | -12.4261
7.2335 | -34.3744 | -12.4357
7.2397 | -34.3747 | -12.4453
7.2458 | -34.3751 | -12.4547
7.2519 | -34.3754 | -12.4643
7.258 | -34.3757 | -12.4741
7.2641 | -34.3761 | -12.4837
7.2702 | -34.3764 | -12.4936
7.2763 | -34.3767 | -12.5032
7.2824 | -34.3771 | -12.513
7.2885 | -34.3774 | -12.5227
7.2946 | -34.3777 | -12.5323
7.3007 | -34.378 | -12.5421
7.3068 | -34.3783 | -12.5519
7.3129 | -34.3787 | -12.5618
7.319 | -34.379 | -12.5714
7.3251 | -34.3793 | -12.5808
7.3312 | -34.3796 | -12.5905
7.3373 | -34.38 | -12.5999
7.3434 | -34.3803 | -12.6094
7.3495 | -34.3807 | -12.6189
7.3556 | -34.381 | -12.6287
7.3617 | -34.3814 | -12.6385
7.3678 | -34.3817 | -12.648
7.3739 | -34.382 | -12.6579
7.3801 | -34.3823 | -12.6678
7.3862 | -34.3827 | -12.6774
7.3923 | -34.383 | -12.6871
7.3984 | -34.3833 | -12.6965
7.4045 | -34.3837 | -12.7061
7.4106 | -34.3841 | -12.7157
7.4167 | -34.3844 | -12.7254
7.4228 | -34.3847 | -12.7352
7.4289 | -34.3851 | -12.7448
7.435 | -34.3854 | -12.7542
7.4411 | -34.3857 | -12.764
7.4472 | -34.3861 | -12.7736
7.4533 | -34.3864 | -12.7831
7.4594 | -34.3868 | -12.793
7.4655 | -34.3871 | -12.8024
7.4716 | -34.3874 | -12.8122
7.4777 | -34.3877 | -12.8218
7.4838 | -34.3881 | -12.8313
7.4899 | -34.3884 | -12.8409
7.496 | -34.3887 | -12.8504
7.5021 | -34.3891 | -12.86
7.5082 | -34.3894 | -12.8696
7.5143 | -34.3898 | -12.8793
7.5204 | -34.3901 | -12.8891
7.5266 | -34.3904 | -12.8986
7.5327 | -34.3908 | -12.9082
7.5388 | -34.3911 | -12.9181
7.5449 | -34.3915 | -12.9279
7.551 | -34.3918 | -12.9374
7.5571 | -34.3921 | -12.9472
7.5632 | -34.3925 | -12.9569
7.5693 | -34.3929 | -12.9664
7.5754 | -34.3932 | -12.976
7.5815 | -34.3935 | -12.9856
7.5876 | -34.3939 | -12.9955
7.5937 | -34.3942 | -13.0053
7.5998 | -34.3946 | -13.0149
7.6059 | -34.395 | -13.0247
7.612 | -34.3953 | -13.0344
7.6181 | -34.3956 | -13.0439
7.6242 | -34.396 | -13.0539
7.6303 | -34.3963 | -13.0637
7.6364 | -34.3966 | -13.0735
7.6425 | -34.397 | -13.0832
7.6486 | -34.3973 | -13.0929
7.6547 | -34.3977 | -13.1024
7.6608 | -34.398 | -13.112
7.667 | -34.3983 | -13.1219
7.6731 | -34.3987 | -13.1316
7.6792 | -34.3991 | -13.1415
7.6853 | -34.3994 | -13.1513
7.6914 | -34.3997 | -13.1611
7.6975 | -34.4001 | -13.1708
7.7036 | -34.4005 | -13.1806
7.7097 | -34.4008 | -13.1904
7.7158 | -34.4012 | -13.2002
7.7219 | -34.4015 | -13.2098
7.728 | -34.4019 | -13.2196
7.7341 | -34.4022 | -13.2294
7.7402 | -34.4025 | -13.2392
7.7463 | -34.4029 | -13.2489
7.7524 | -34.4032 | -13.2589
7.7585 | -34.4036 | -13.2686
7.7646 | -34.4039 | -13.2783
7.7707 | -34.4043 | -13.2882
7.7768 | -34.4046 | -13.2981
7.7829 | -34.4049 | -13.3081
7.789 | -34.4053 | -13.3181
7.7951 | -34.4056 | -13.3278
7.8012 | -34.4059 | -13.3375
7.8073 | -34.4063 | -13.3475
7.8135 | -34.4067 | -13.3572
7.8196 | -34.4071 | -13.3672
7.8257 | -34.4074 | -13.3772
7.8318 | -34.4077 | -13.387
7.8379 | -34.4081 | -13.3967
7.844 | -34.4084 | -13.4064
7.8501 | -34.4088 | -13.4163
7.8562 | -34.4091 | -13.4263
7.8623 | -34.4095 | -13.436
7.8684 | -34.4098 | -13.4458
7.8745 | -34.4102 | -13.4556
7.8806 | -34.4105 | -13.4656
7.8867 | -34.4109 | -13.4752
7.8928 | -34.4112 | -13.4847
7.8989 | -34.4116 | -13.4945
7.905 | -34.4119 | -13.5044
7.9111 | -34.4123 | -13.5142
7.9172 | -34.4126 | -13.524
7.9233 | -34.4129 | -13.5339
7.9294 | -34.4133 | -13.5437
7.9355 | -34.4136 | -13.5533
7.9416 | -34.414 | -13.5632
7.9477 | -34.4143 | -13.5728
7.9539 | -34.4147 | -13.5826
7.96 | -34.415 | -13.5925
7.9661 | -34.4154 | -13.6024
7.9722 | -34.4158 | -13.612
7.9783 | -34.4161 | -13.6221
7.9844 | -34.4165 | -13.6318
7.9905 | -34.4168 | -13.6418
7.9966 | -34.4172 | -13.6518
8.0027 | -34.4175 | -13.6616
8.0088 | -34.4178 | -13.6714
8.0149 | -34.4182 | -13.6812
8.021 | -34.4185 | -13.6911
8.0271 | -34.4189 | -13.7012
8.0332 | -34.4192 | -13.7109
8.0393 | -34.4196 | -13.7209
8.0454 | -34.42 | -13.731
8.0515 | -34.4203 | -13.7408
8.0576 | -34.4207 | -13.7509
8.0637 | -34.421 | -13.7606
8.0698 | -34.4214 | -13.7703
8.0759 | -34.4217 | -13.78
8.082 | -34.422 | -13.7897
8.0881 | -34.4224 | -13.7996
8.0942 | -34.4228 | -13.8095
8.1004 | -34.4231 | -13.8194
8.1065 | -34.4235 | -13.8291
8.1126 | -34.4238 | -13.8389
8.1187 | -34.4241 | -13.8489
8.1248 | -34.4245 | -13.8585
8.1309 | -34.4249 | -13.8682
8.137 | -34.4252 | -13.8778
8.1431 | -34.4255 | -13.8879
8.1492 | -34.4259 | -13.8976
8.1553 | -34.4262 | -13.9075
8.1614 | -34.4266 | -13.9175
8.1675 | -34.4269 | -13.9273
8.1736 | -34.4272 | -13.9373
8.1797 | -34.4276 | -13.9472
8.1858 | -34.428 | -13.9571
8.1919 | -34.4283 | -13.9669
8.198 | -34.4287 | -13.9766
8.2041 | -34.429 | -13.9864
8.2102 | -34.4294 | -13.9962
8.2163 | -34.4297 | -14.0059
8.2224 | -34.4301 | -14.0157
8.2285 | -34.4304 | -14.0253
8.2346 | -34.4308 | -14.035
8.2408 | -34.4311 | -14.0447
8.2469 | -34.4315 | -14.0546
8.253 | -34.4319 | -14.0643
8.2591 | -34.4322 | -14.0741
8.2652 | -34.4326 | -14.0839
8.2713 | -34.433 | -14.0938
8.2774 | -34.4333 | -14.1037
8.2835 | -34.4337 | -14.1138
8.2896 | -34.434 | -14.1234
8.2957 | -34.4344 | -14.1333
8.3018 | -34.4348 | -14.1432
8.3079 | -34.4351 | -14.1528
8.314 | -34.4355 | -14.1626
8.3201 | -34.4359 | -14.1723
8.3262 | -34.4362 | -14.1822
8.3323 | -34.4365 | -14.1919
8.3384 | -34.4369 | -14.2019
8.3445 | -34.4372 | -14.2116
8.3506 | -34.4376 | -14.2214
8.3567 | -34.438 | -14.2313
8.3628 | -34.4383 | -14.2411
8.3689 | -34.4387 | -14.2507
8.375 | -34.439 | -14.2607
8.3812 | -34.4394 | -14.2704
8.3873 | -34.4397 | -14.2801
8.3934 | -34.4401 | -14.29
8.3995 | -34.4404 | -14.2997
8.4056 | -34.4408 | -14.3095
8.4117 | -34.4411 | -14.3193
8.4178 | -34.4415 | -14.3292
8.4239 | -34.4418 | -14.3388
8.43 | -34.4422 | -14.3486
8.4361 | -34.4425 | -14.3585
8.4422 | -34.4429 | -14.3681
8.4483 | -34.4432 | -14.3778
8.4544 | -34.4436 | -14.3876
8.4605 | -34.4439 | -14.3972
8.4666 | -34.4443 | -14.4072
8.4727 | -34.4446 | -14.417
8.4788 | -34.445 | -14.4267
8.4849 | -34.4453 | -14.4365
8.491 | -34.4457 | -14.4463
8.4971 | -34.446 | -14.4558
8.5032 | -34.4464 | -14.4655
8.5093 | -34.4467 | -14.4751
8.5154 | -34.4471 | -14.4849
8.5215 | -34.4475 | -14.4946
8.5277 | -34.4478 | -14.5044
8.5338 | -34.4482 | -14.514
8.5399 | -34.4485 | -14.5238
8.546 | -34.4489 | -14.5333
8.5521 | -34.4493 | -14.5431
8.5582 | -34.4496 | -14.5528
8.5643 | -34.4499 | -14.5623
8.5704 | -34.4503 | -14.5717
8.5765 | -34.4507 | -14.5816
8.5826 | -34.451 | -14.591
8.5887 | -34.4514 | -14.6007
8.5948 | -34.4518 | -14.6105
8.6009 | -34.4521 | -14.62
8.607 | -34.4525 | -14.6298
8.6131 | -34.4528 | -14.6396
8.6192 | -34.4532 | -14.6491
8.6253 | -34.4536 | -14.6587
8.6314 | -34.4539 | -14.6683
8.6375 | -34.4542 | -14.678
8.6436 | -34.4546 | -14.6877
8.6497 | -34.4549 | -14.6973
8.6558 | -34.4553 | -14.7069
8.6619 | -34.4556 | -14.7166
8.6681 | -34.456 | -14.7264
8.6742 | -34.4564 | -14.7361
8.6803 | -34.4567 | -14.7457
8.6864 | -34.4571 | -14.7551
8.6925 | -34.4574 | -14.7646
8.6986 | -34.4578 | -14.7743
8.7047 | -34.4581 | -14.7841
8.7108 | -34.4584 | -14.7938
8.7169 | -34.4588 | -14.8035
8.723 | -34.4592 | -14.8129
8.7291 | -34.4595 | -14.8224
8.7352 | -34.4599 | -14.8322
8.7413 | -34.4602 | -14.8419
8.7474 | -34.4606 | -14.8513
8.7535 | -34.461 | -14.861
8.7596 | -34.4613 | -14.8706
8.7657 | -34.4617 | -14.8803
8.7718 | -34.462 | -14.89
8.7779 | -34.4624 | -14.8997
8.784 | -34.4627 | -14.9094
8.7901 | -34.4631 | -14.9189
8.7962 | -34.4635 | -14.9283
8.8023 | -34.4638 | -14.9379
8.8084 | -34.4642 | -14.9475
8.8146 | -34.4646 | -14.957
8.8207 | -34.4649 | -14.9665
8.8268 | -34.4653 | -14.9757
8.8329 | -34.4656 | -14.9853
8.839 | -34.466 | -14.9949
8.8451 | -34.4663 | -15.0045
8.8512 | -34.4667 | -15.014
8.8573 | -34.4671 | -15.0236
8.8634 | -34.4674 | -15.033
8.8695 | -34.4678 | -15.0428
8.8756 | -34.4681 | -15.0523
8.8817 | -34.4685 | -15.0618
8.8878 | -34.4688 | -15.0716
8.8939 | -34.4692 | -15.0811
8.9 | -34.4696 | -15.0906
8.9061 | -34.4699 | -15.1001
8.9122 | -34.4703 | -15.1097
8.9183 | -34.4707 | -15.1193
8.9244 | -34.4711 | -15.129
8.9305 | -34.4715 | -15.1383
8.9366 | -34.4718 | -15.1478
8.9427 | -34.4722 | -15.1574
8.9488 | -34.4725 | -15.1668
8.955 | -34.4729 | -15.1763
8.9611 | -34.4733 | -15.186
8.9672 | -34.4737 | -15.1954
8.9733 | -34.474 | -15.2048
8.9794 | -34.4744 | -15.2142
8.9855 | -34.4748 | -15.2238
8.9916 | -34.4751 | -15.2335
8.9977 | -34.4755 | -15.243
9.0038 | -34.4759 | -15.2523
9.0099 | -34.4763 | -15.2619
9.016 | -34.4766 | -15.2715
9.0221 | -34.477 | -15.2809
9.0282 | -34.4774 | -15.2903
9.0343 | -34.4778 | -15.2999
9.0404 | -34.4781 | -15.3094
9.0465 | -34.4785 | -15.319
9.0526 | -34.4789 | -15.3284
9.0587 | -34.4793 | -15.338
9.0648 | -34.4797 | -15.3473
9.0709 | -34.48 | -15.3567
9.077 | -34.4805 | -15.3663
9.0831 | -34.4808 | -15.376
9.0892 | -34.4812 | -15.3856
9.0953 | -34.4816 | -15.3951
9.1015 | -34.482 | -15.4045
9.1076 | -34.4824 | -15.4139
9.1137 | -34.4827 | -15.4234
9.1198 | -34.4831 | -15.4328
9.1259 | -34.4835 | -15.4423
9.132 | -34.4839 | -15.4519
9.1381 | -34.4843 | -15.4612
9.1442 | -34.4847 | -15.4705
9.1503 | -34.485 | -15.4801
9.1564 | -34.4854 | -15.4895
9.1625 | -34.4858 | -15.499
9.1686 | -34.4862 | -15.5083
9.1747 | -34.4866 | -15.5177
9.1808 | -34.487 | -15.5271
9.1869 | -34.4874 | -15.5365
9.193 | -34.4878 | -15.546
9.1991 | -34.4881 | -15.5555
9.2052 | -34.4885 | -15.565
9.2113 | -34.4889 | -15.5744
9.2174 | -34.4894 | -15.5838
9.2235 | -34.4898 | -15.5933
9.2296 | -34.4902 | -15.6027
9.2357 | -34.4906 | -15.6123
9.2419 | -34.491 | -15.6217
9.248 | -34.4914 | -15.6311
9.2541 | -34.4918 | -15.6406
9.2602 | -34.4922 | -15.6502
9.2663 | -34.4926 | -15.6594
9.2724 | -34.493 | -15.6688
9.2785 | -34.4933 | -15.6782
9.2846 | -34.4938 | -15.6875
9.2907 | -34.4941 | -15.6969
9.2968 | -34.4945 | -15.7064
9.3029 | -34.4949 | -15.716
9.309 | -34.4953 | -15.7254
9.3151 | -34.4957 | -15.7348
9.3212 | -34.4961 | -15.7441
9.3273 | -34.4965 | -15.7535
9.3334 | -34.4969 | -15.7628
9.3395 | -34.4973 | -15.7721
9.3456 | -34.4977 | -15.7814
9.3517 | -34.498 | -15.7908
9.3578 | -34.4984 | -15.8003
9.3639 | -34.4989 | -15.8096
9.37 | -34.4992 | -15.8191
9.3761 | -34.4996 | -15.8287
9.3822 | -34.5 | -15.8382
9.3884 | -34.5004 | -15.8473
9.3945 | -34.5008 | -15.8568
9.4006 | -34.5013 | -15.8662
9.4067 | -34.5016 | -15.8756
9.4128 | -34.5021 | -15.8852
9.4189 | -34.5025 | -15.8948
9.425 | -34.5029 | -15.9041
9.4311 | -34.5033 | -15.9135
9.4372 | -34.5037 | -15.9228
9.4433 | -34.5041 | -15.9322
9.4494 | -34.5045 | -15.9414
9.4555 | -34.5049 | -15.951
9.4616 | -34.5053 | -15.9604
9.4677 | -34.5057 | -15.9697
9.4738 | -34.5061 | -15.9791
9.4799 | -34.5065 | -15.9885
9.486 | -34.5069 | -15.998
9.4921 | -34.5073 | -16.0074
9.4982 | -34.5077 | -16.0165
9.5043 | -34.5081 | -16.0259
9.5104 | -34.5085 | -16.0351
9.5165 | -34.5089 | -16.0445
9.5226 | -34.5093 | -16.054
9.5288 | -34.5097 | -16.0631
9.5349 | -34.5101 | -16.0725
9.541 | -34.5105 | -16.0817
9.5471 | -34.5109 | -16.091
9.5532 | -34.5114 | -16.1004
9.5593 | -34.5118 | -16.1098
9.5654 | -34.5122 | -16.1191
9.5715 | -34.5126 | -16.1284
9.5776 | -34.513 | -16.138
9.5837 | -34.5133 | -16.1474
9.5898 | -34.5137 | -16.1568
9.5959 | -34.5141 | -16.1662
9.602 | -34.5145 | -16.1754
9.6081 | -34.5149 | -16.1847
9.6142 | -34.5154 | -16.1941
9.6203 | -34.5157 | -16.2034
9.6264 | -34.5161 | -16.2128
9.6325 | -34.5165 | -16.2222
9.6386 | -34.5169 | -16.2314
9.6447 | -34.5173 | -16.2407
9.6508 | -34.5177 | -16.25
9.6569 | -34.5181 | -16.2594
9.663 | -34.5185 | -16.2688
9.6691 | -34.5189 | -16.2781
9.6753 | -34.5193 | -16.2874
9.6814 | -34.5197 | -16.2966
9.6875 | -34.5201 | -16.3058
9.6936 | -34.5205 | -16.315
9.6997 | -34.521 | -16.3244
9.7058 | -34.5214 | -16.3337
9.7119 | -34.5218 | -16.3431
9.718 | -34.5222 | -16.3526
9.7241 | -34.5226 | -16.3618
9.7302 | -34.523 | -16.3711
9.7363 | -34.5234 | -16.3804
9.7424 | -34.5238 | -16.3896
9.7485 | -34.5242 | -16.399
9.7546 | -34.5246 | -16.4082
9.7607 | -34.525 | -16.4178
9.7668 | -34.5255 | -16.4272
9.7729 | -34.5259 | -16.4364
9.779 | -34.5263 | -16.4456
9.7851 | -34.5267 | -16.455
9.7912 | -34.5271 | -16.4643
9.7973 | -34.5276 | -16.4738
9.8034 | -34.528 | -16.4829
9.8095 | -34.5284 | -16.4924
9.8157 | -34.5288 | -16.5018
9.8218 | -34.5292 | -16.5113
9.8279 | -34.5296 | -16.5205
9.834 | -34.53 | -16.53
9.8401 | -34.5304 | -16.5392
9.8462 | -34.5308 | -16.5482
9.8523 | -34.5312 | -16.5575
9.8584 | -34.5316 | -16.5668
9.8645 | -34.532 | -16.5762
9.8706 | -34.5324 | -16.5855
9.8767 | -34.5328 | -16.5947
9.8828 | -34.5332 | -16.6039
9.8889 | -34.5336 | -16.6131
9.895 | -34.534 | -16.6224
9.9011 | -34.5344 | -16.6317
9.9072 | -34.5348 | -16.641
9.9133 | -34.5352 | -16.6503
9.9194 | -34.5356 | -16.6595
9.9255 | -34.536 | -16.6687
9.9316 | -34.5364 | -16.6779
9.9377 | -34.5368 | -16.6872
9.9438 | -34.5372 | -16.6965
9.9499 | -34.5376 | -16.7061
9.956 | -34.538 | -16.7153
9.9622 | -34.5384 | -16.7247
9.9683 | -34.5388 | -16.734
9.9744 | -34.5392 | -16.7434
9.9805 | -34.5396 | -16.7527
9.9866 | -34.54 | -16.7619
9.9927 | -34.5404 | -16.771
9.9988 | -34.5408 | -16.7804
10.0049 | -34.5412 | -16.7897
10.011 | -34.5416 | -16.799
10.0171 | -34.542 | -16.8082
10.0232 | -34.5424 | -16.8174
10.0293 | -34.5429 | -16.8267
10.0354 | -34.5433 | -16.8361
10.0415 | -34.5437 | -16.8454
10.0476 | -34.5441 | -16.8548
10.0537 | -34.5445 | -16.8643
10.0598 | -34.5449 | -16.8736
10.0659 | -34.5453 | -16.8829
10.072 | -34.5457 | -16.892
10.0781 | -34.5461 | -16.9012
10.0842 | -34.5465 | -16.9104
10.0903 | -34.5469 | -16.9197
10.0964 | -34.5474 | -16.9288
10.1026 | -34.5477 | -16.9382
10.1087 | -34.5481 | -16.9475
10.1148 | -34.5485 | -16.9568
10.1209 | -34.5489 | -16.9659
10.127 | -34.5493 | -16.9751
10.1331 | -34.5498 | -16.9843
10.1392 | -34.5502 | -16.9934
10.1453 | -34.5506 | -17.0027
10.1514 | -34.551 | -17.012
10.1575 | -34.5513 | -17.0214
10.1636 | -34.5518 | -17.0304
10.1697 | -34.5522 | -17.0397
10.1758 | -34.5526 | -17.0488
10.1819 | -34.553 | -17.0579
10.188 | -34.5533 | -17.0672
10.1941 | -34.5537 | -17.0765
10.2002 | -34.5541 | -17.0858
10.2063 | -34.5545 | -17.095
10.2124 | -34.5549 | -17.1043
10.2185 | -34.5553 | -17.1136
10.2246 | -34.5557 | -17.1227
10.2307 | -34.5561 | -17.132
10.2368 | -34.5565 | -17.1408
10.2429 | -34.5569 | -17.15
10.2491 | -34.5573 | -17.1592
10.2552 | -34.5577 | -17.1686
10.2613 | -34.5581 | -17.178
10.2674 | -34.5585 | -17.1874
10.2735 | -34.5589 | -17.1968
10.2796 | -34.5593 | -17.2061
10.2857 | -34.5597 | -17.2154
10.2918 | -34.5601 | -17.2249
10.2979 | -34.5605 | -17.2341
10.304 | -34.5609 | -17.2432
10.3101 | -34.5613 | -17.2526
10.3162 | -34.5617 | -17.262
10.3223 | -34.5621 | -17.271
10.3284 | -34.5625 | -17.2802
10.3345 | -34.5629 | -17.2893
10.3406 | -34.5633 | -17.2985
10.3467 | -34.5636 | -17.3076
10.3528 | -34.564 | -17.3166
10.3589 | -34.5644 | -17.3258
10.365 | -34.5648 | -17.3352
10.3711 | -34.5652 | -17.3444
10.3772 | -34.5656 | -17.3538
10.3833 | -34.566 | -17.363
10.3895 | -34.5664 | -17.3722
10.3956 | -34.5668 | -17.3814
10.4017 | -34.5673 | -17.3905
10.4078 | -34.5676 | -17.3996
10.4139 | -34.5681 | -17.409
10.42 | -34.5685 | -17.4185
10.4261 | -34.5689 | -17.4277
10.4322 | -34.5692 | -17.437
10.4383 | -34.5696 | -17.4462
10.4444 | -34.57 | -17.4552
10.4505 | -34.5704 | -17.4644
10.4566 | -34.5708 | -17.4736
10.4627 | -34.5712 | -17.4827
10.4688 | -34.5716 | -17.4918
10.4749 | -34.572 | -17.5011
10.481 | -34.5724 | -17.5105
10.4871 | -34.5728 | -17.5196
10.4932 | -34.5732 | -17.5286
10.4993 | -34.5736 | -17.538
10.5054 | -34.574 | -17.5473
10.5115 | -34.5744 | -17.5564
10.5176 | -34.5748 | -17.5654
10.5237 | -34.5752 | -17.5749
10.5298 | -34.5756 | -17.584
10.536 | -34.576 | -17.5932
10.5421 | -34.5764 | -17.6022
10.5482 | -34.5768 | -17.6115
10.5543 | -34.5772 | -17.6206
10.5604 | -34.5776 | -17.6296
10.5665 | -34.578 | -17.6386
10.5726 | -34.5784 | -17.6478
10.5787 | -34.5788 | -17.6569
10.5848 | -34.5791 | -17.666
10.5909 | -34.5796 | -17.6753
10.597 | -34.5799 | -17.6846
10.6031 | -34.5803 | -17.694
10.6092 | -34.5807 | -17.7032
10.6153 | -34.5811 | -17.7126
10.6214 | -34.5815 | -17.7217
10.6275 | -34.5819 | -17.731
10.6336 | -34.5823 | -17.7401
10.6397 | -34.5827 | -17.7494
10.6458 | -34.5831 | -17.7584
10.6519 | -34.5835 | -17.7678
10.658 | -34.5839 | -17.7771
10.6641 | -34.5843 | -17.7864
10.6702 | -34.5847 | -17.7956
10.6764 | -34.5851 | -17.8049
10.6825 | -34.5855 | -17.8139
10.6886 | -34.5859 | -17.8229
10.6947 | -34.5863 | -17.832
10.7008 | -34.5867 | -17.8411
10.7069 | -34.5871 | -17.8504
10.713 | -34.5875 | -17.8595
10.7191 | -34.5879 | -17.8687
10.7252 | -34.5882 | -17.8776
10.7313 | -34.5886 | -17.8867
10.7374 | -34.589 | -17.896
10.7435 | -34.5894 | -17.9051
10.7496 | -34.5897 | -17.9141
10.7557 | -34.5901 | -17.9235
10.7618 | -34.5905 | -17.9324
10.7679 | -34.5909 | -17.9415
10.774 | -34.5914 | -17.9508
10.7801 | -34.5918 | -17.9599
10.7862 | -34.5922 | -17.969
10.7923 | -34.5925 | -17.9782
10.7984 | -34.5929 | -17.9872
10.8045 | -34.5934 | -17.9962
10.8106 | -34.5937 | -18.0055
10.8168 | -34.5941 | -18.0147
10.8229 | -34.5945 | -18.0239
10.829 | -34.5949 | -18.0331
10.8351 | -34.5953 | -18.042
10.8412 | -34.5957 | -18.0511
10.8473 | -34.5961 | -18.0601
10.8534 | -34.5965 | -18.0692
10.8595 | -34.5969 | -18.0783
10.8656 | -34.5974 | -18.0875
10.8717 | -34.5978 | -18.0967
10.8778 | -34.5982 | -18.106
10.8839 | -34.5986 | -18.1151
10.89 | -34.599 | -18.1243
10.8961 | -34.5994 | -18.1334
10.9022 | -34.5998 | -18.1427
10.9083 | -34.6002 | -18.1518
10.9144 | -34.6006 | -18.1608
10.9205 | -34.601 | -18.1699
10.9266 | -34.6014 | -18.1789
10.9327 | -34.6018 | -18.188
10.9388 | -34.6022 | -18.1971
10.9449 | -34.6025 | -18.2061
10.951 | -34.6029 | -18.2153
10.9571 | -34.6033 | -18.2245
10.9633 | -34.6037 | -18.2339
10.9694 | -34.6041 | -18.2431
10.9755 | -34.6045 | -18.2521
10.9816 | -34.6049 | -18.2613
10.9877 | -34.6053 | -18.2703
10.9938 | -34.6057 | -18.2795
10.9999 | -34.6062 | -18.2885
11.006 | -34.6066 | -18.2975
11.0121 | -34.607 | -18.3067
11.0182 | -34.6074 | -18.3157
11.0243 | -34.6078 | -18.3249
11.0304 | -34.6082 | -18.334
11.0365 | -34.6086 | -18.3431
11.0426 | -34.609 | -18.352
11.0487 | -34.6094 | -18.3612
11.0548 | -34.6098 | -18.3702
11.0609 | -34.6102 | -18.3793
11.067 | -34.6106 | -18.3884
11.0731 | -34.6111 | -18.3975
11.0792 | -34.6115 | -18.4065
11.0853 | -34.6119 | -18.4157
11.0914 | -34.6123 | -18.4249
11.0975 | -34.6127 | -18.434
11.1037 | -34.6131 | -18.443
11.1098 | -34.6135 | -18.4522
11.1159 | -34.6139 | -18.461
11.122 | -34.6143 | -18.47
11.1281 | -34.6147 | -18.4791
11.1342 | -34.6152 | -18.4882
11.1403 | -34.6156 | -18.4973
11.1464 | -34.616 | -18.5065
11.1525 | -34.6164 | -18.5153
11.1586 | -34.6168 | -18.5244
11.1647 | -34.6172 | -18.5334
11.1708 | -34.6176 | -18.5425
11.1769 | -34.618 | -18.5516
11.183 | -34.6184 | -18.5605
11.1891 | -34.6188 | -18.5695
11.1952 | -34.6192 | -18.5785
11.2013 | -34.6196 | -18.5875
11.2074 | -34.62 | -18.5964
11.2135 | -34.6205 | -18.6057
11.2196 | -34.6209 | -18.6148
11.2257 | -34.6213 | -18.624
11.2318 | -34.6218 | -18.6331
11.2379 | -34.6222 | -18.6422
11.244 | -34.6227 | -18.6513
11.2502 | -34.6231 | -18.6603
11.2563 | -34.6235 | -18.6695
11.2624 | -34.624 | -18.6784
11.2685 | -34.6244 | -18.6874
11.2746 | -34.6248 | -18.6964
11.2807 | -34.6252 | -18.7056
11.2868 | -34.6256 | -18.7148
11.2929 | -34.626 | -18.7235
11.299 | -34.6265 | -18.7326
11.3051 | -34.6269 | -18.7419
11.3112 | -34.6273 | -18.7509
11.3173 | -34.6277 | -18.7598
11.3234 | -34.6282 | -18.7687
11.3295 | -34.6286 | -18.7778
11.3356 | -34.629 | -18.7869
11.3417 | -34.6294 | -18.7958
11.3478 | -34.6298 | -18.805
11.3539 | -34.6302 | -18.8142
11.36 | -34.6306 | -18.8231
11.3661 | -34.6311 | -18.832
11.3722 | -34.6315 | -18.8408
11.3783 | -34.6319 | -18.85
11.3844 | -34.6324 | -18.8591
11.3906 | -34.6328 | -18.8682
11.3967 | -34.6332 | -18.877
11.4028 | -34.6336 | -18.8862
11.4089 | -34.6341 | -18.8951
11.415 | -34.6345 | -18.904
11.4211 | -34.6349 | -18.9131
11.4272 | -34.6353 | -18.9221
11.4333 | -34.6357 | -18.931
11.4394 | -34.6361 | -18.94
11.4455 | -34.6365 | -18.9489
11.4516 | -34.6369 | -18.9579
11.4577 | -34.6373 | -18.9669
11.4638 | -34.6378 | -18.9756
11.4699 | -34.6382 | -18.9847
11.476 | -34.6386 | -18.9939
11.4821 | -34.639 | -19.0028
11.4882 | -34.6395 | -19.0121
11.4943 | -34.6399 | -19.0212
11.5004 | -34.6404 | -19.0297
11.5065 | -34.6408 | -19.0387
11.5126 | -34.6412 | -19.0477
11.5187 | -34.6416 | -19.0569
11.5248 | -34.642 | -19.066
11.5309 | -34.6425 | -19.075
11.5371 | -34.6429 | -19.0837
11.5432 | -34.6433 | -19.0928
11.5493 | -34.6438 | -19.102
11.5554 | -34.6442 | -19.1109
11.5615 | -34.6446 | -19.12
11.5676 | -34.6451 | -19.129
11.5737 | -34.6455 | -19.1378
11.5798 | -34.6459 | -19.1465
11.5859 | -34.6464 | -19.1555
11.592 | -34.6468 | -19.1644
11.5981 | -34.6472 | -19.1733
11.6042 | -34.6476 | -19.1821
11.6103 | -34.6481 | -19.1912
11.6164 | -34.6485 | -19.1999
11.6225 | -34.6489 | -19.2086
11.6286 | -34.6493 | -19.2177
11.6347 | -34.6498 | -19.2266
11.6408 | -34.6502 | -19.2354
11.6469 | -34.6507 | -19.2443
11.653 | -34.6511 | -19.2531
11.6591 | -34.6515 | -19.262
11.6652 | -34.652 | -19.2709
11.6713 | -34.6524 | -19.2797
11.6775 | -34.6529 | -19.2886
11.6836 | -34.6533 | -19.2975
11.6897 | -34.6538 | -19.3065
11.6958 | -34.6542 | -19.3155
11.7019 | -34.6547 | -19.3245
11.708 | -34.6551 | -19.3334
11.7141 | -34.6556 | -19.3421
11.7202 | -34.656 | -19.3511
11.7263 | -34.6564 | -19.3601
11.7324 | -34.6569 | -19.3692
11.7385 | -34.6573 | -19.3782
11.7446 | -34.6577 | -19.387
11.7507 | -34.6582 | -19.3956
11.7568 | -34.6586 | -19.4046
11.7629 | -34.659 | -19.4136
11.769 | -34.6595 | -19.4224
11.7751 | -34.6599 | -19.4312
11.7812 | -34.6604 | -19.4397
11.7873 | -34.6608 | -19.4485
11.7934 | -34.6613 | -19.4574
11.7995 | -34.6617 | -19.4663
11.8056 | -34.6622 | -19.4751
11.8117 | -34.6626 | -19.4842
11.8178 | -34.6631 | -19.4928
11.824 | -34.6635 | -19.5016
11.8301 | -34.6639 | -19.5104
11.8362 | -34.6644 | -19.5191
11.8423 | -34.6648 | -19.528
11.8484 | -34.6652 | -19.5368
11.8545 | -34.6657 | -19.5454
11.8606 | -34.6661 | -19.5542
11.8667 | -34.6666 | -19.563
11.8728 | -34.667 | -19.5718
11.8789 | -34.6674 | -19.5807
11.885 | -34.6679 | -19.5896
11.8911 | -34.6683 | -19.5982
11.8972 | -34.6688 | -19.6071
11.9033 | -34.6692 | -19.6158
11.9094 | -34.6696 | -19.6245
11.9155 | -34.6701 | -19.6333
11.9216 | -34.6706 | -19.6419
11.9277 | -34.671 | -19.6507
11.9338 | -34.6714 | -19.6596
11.9399 | -34.6719 | -19.6685
11.946 | -34.6723 | -19.6772
11.9521 | -34.6727 | -19.686
11.9582 | -34.6732 | -19.6947
11.9644 | -34.6737 | -19.7033
11.9705 | -34.6741 | -19.7121
11.9766 | -34.6746 | -19.7209
11.9827 | -34.6751 | -19.7296
11.9888 | -34.6755 | -19.7383
11.9949 | -34.676 | -19.7468
12.001 | -34.6765 | -19.7556
12.0071 | -34.6769 | -19.7644
12.0132 | -34.6773 | -19.773
12.0193 | -34.6778 | -19.7818
12.0254 | -34.6782 | -19.7906
12.0315 | -34.6787 | -19.799
12.0376 | -34.6791 | -19.8077
12.0437 | -34.6795 | -19.8163
12.0498 | -34.68 | -19.825
12.0559 | -34.6804 | -19.8337
12.062 | -34.6809 | -19.8422
12.0681 | -34.6814 | -19.851
12.0742 | -34.6818 | -19.8596
12.0803 | -34.6823 | -19.8685
12.0864 | -34.6827 | -19.8773
12.0925 | -34.6832 | -19.886
12.0986 | -34.6836 | -19.8945
12.1047 | -34.6841 | -19.9032
12.1109 | -34.6846 | -19.9118
12.117 | -34.685 | -19.9205
12.1231 | -34.6855 | -19.9291
12.1292 | -34.6859 | -19.9378
12.1353 | -34.6864 | -19.9463
12.1414 | -34.6868 | -19.955
12.1475 | -34.6873 | -19.9638
12.1536 | -34.6877 | -19.9725
12.1597 | -34.6881 | -19.9812
12.1658 | -34.6886 | -19.9899
12.1719 | -34.689 | -19.9983
12.178 | -34.6895 | -20.007
12.1841 | -34.69 | -20.0158
12.1902 | -34.6904 | -20.0244
12.1963 | -34.6909 | -20.033
12.2024 | -34.6914 | -20.0415
12.2085 | -34.6918 | -20.0502
12.2146 | -34.6923 | -20.0589
12.2207 | -34.6927 | -20.0674
12.2268 | -34.6932 | -20.076
12.2329 | -34.6937 | -20.0847
12.239 | -34.6941 | -20.093
12.2451 | -34.6946 | -20.1017
12.2513 | -34.695 | -20.1104
12.2574 | -34.6955 | -20.1191
12.2635 | -34.696 | -20.1278
12.2696 | -34.6964 | -20.1363
12.2757 | -34.6969 | -20.1446
12.2818 | -34.6974 | -20.1534
12.2879 | -34.6978 | -20.162
12.294 | -34.6983 | -20.1704
12.3001 | -34.6988 | -20.1791
12.3062 | -34.6992 | -20.1877
12.3123 | -34.6997 | -20.196
12.3184 | -34.7001 | -20.2045
12.3245 | -34.7006 | -20.213
12.3306 | -34.7011 | -20.2216
12.3367 | -34.7015 | -20.2301
12.3428 | -34.702 | -20.2387
12.3489 | -34.7024 | -20.2469
12.355 | -34.7029 | -20.2556
12.3611 | -34.7034 | -20.2642
12.3672 | -34.7038 | -20.2727
12.3733 | -34.7043 | -20.2812
12.3794 | -34.7047 | -20.2895
12.3855 | -34.7052 | -20.2979
12.3916 | -34.7057 | -20.3065
12.3978 | -34.7061 | -20.3152
12.4039 | -34.7066 | -20.3238
12.41 | -34.7071 | -20.3323
12.4161 | -34.7075 | -20.3405
12.4222 | -34.708 | -20.3491
12.4283 | -34.7084 | -20.3576
12.4344 | -34.7089 | -20.3662
12.4405 | -34.7094 | -20.3748
12.4466 | -34.7098 | -20.3832
12.4527 | -34.7103 | -20.3915
12.4588 | -34.7107 | -20.4
12.4649 | -34.7112 | -20.4086
12.471 | -34.7117 | -20.4172
12.4771 | -34.7121 | -20.4257
12.4832 | -34.7126 | -20.4342
12.4893 | -34.713 | -20.4425
12.4954 | -34.7135 | -20.4509
12.5015 | -34.714 | -20.4594
12.5076 | -34.7144 | -20.4679
12.5137 | -34.7149 | -20.4764
12.5198 | -34.7154 | -20.4847
12.5259 | -34.7158 | -20.4932
12.532 | -34.7162 | -20.5017
12.5382 | -34.7167 | -20.5103
12.5443 | -34.7172 | -20.5187
12.5504 | -34.7176 | -20.5273
12.5565 | -34.7181 | -20.5355
12.5626 | -34.7186 | -20.5441
12.5687 | -34.719 | -20.5526
12.5748 | -34.7195 | -20.5611
12.5809 | -34.7199 | -20.5695
12.587 | -34.7204 | -20.5781
12.5931 | -34.721 | -20.5862
12.5992 | -34.7215 | -20.5946
12.6053 | -34.7219 | -20.603
12.6114 | -34.7224 | -20.6113
12.6175 | -34.7229 | -20.6199
12.6236 | -34.7233 | -20.6284
12.6297 | -34.7238 | -20.6367
12.6358 | -34.7242 | -20.6449
12.6419 | -34.7246 | -20.6534
12.648 | -34.7251 | -20.662
12.6541 | -34.7256 | -20.6704
12.6602 | -34.726 | -20.6784
12.6663 | -34.7265 | -20.6868
12.6724 | -34.727 | -20.6951
12.6785 | -34.7274 | -20.7035
12.6847 | -34.7279 | -20.7118
12.6908 | -34.7283 | -20.7203
12.6969 | -34.7288 | -20.7285
12.703 | -34.7293 | -20.737
12.7091 | -34.7297 | -20.7454
12.7152 | -34.7302 | -20.7537
12.7213 | -34.7306 | -20.7622
12.7274 | -34.7311 | -20.7705
12.7335 | -34.7316 | -20.7786
12.7396 | -34.7321 | -20.7869
12.7457 | -34.7325 | -20.7953
12.7518 | -34.733 | -20.8037
12.7579 | -34.7334 | -20.8122
12.764 | -34.7339 | -20.8206
12.7701 | -34.7343 | -20.829
12.7762 | -34.7348 | -20.8373
12.7823 | -34.7353 | -20.8456
12.7884 | -34.7358 | -20.8541
12.7945 | -34.7362 | -20.8627
12.8006 | -34.7367 | -20.8709
12.8067 | -34.7371 | -20.8792
12.8128 | -34.7376 | -20.8874
12.8189 | -34.738 | -20.8957
12.8251 | -34.7385 | -20.904
12.8312 | -34.7389 | -20.9124
12.8373 | -34.7394 | -20.9207
12.8434 | -34.7399 | -20.929
12.8495 | -34.7403 | -20.9373
12.8556 | -34.7408 | -20.9457
12.8617 | -34.7413 | -20.9541
12.8678 | -34.7417 | -20.9626
12.8739 | -34.7422 | -20.9706
12.88 | -34.7426 | -20.9789
12.8861 | -34.743 | -20.9872
12.8922 | -34.7435 | -20.9956
12.8983 | -34.744 | -21.0039
12.9044 | -34.7445 | -21.012
12.9105 | -34.7449 | -21.0202
12.9166 | -34.7454 | -21.0284
12.9227 | -34.7458 | -21.0366
12.9288 | -34.7463 | -21.045
12.9349 | -34.7467 | -21.0534
12.941 | -34.7472 | -21.0617
12.9471 | -34.7477 | -21.0701
12.9532 | -34.7482 | -21.0783
12.9593 | -34.7486 | -21.0868
12.9654 | -34.749 | -21.0952
12.9716 | -34.7495 | -21.1036
12.9777 | -34.75 | -21.1119
12.9838 | -34.7505 | -21.1202
12.9899 | -34.7509 | -21.1286
12.996 | -34.7514 | -21.1369
13.0021 | -34.7518 | -21.1452
13.0082 | -34.7523 | -21.1534
13.0143 | -34.7528 | -21.1612
13.0204 | -34.7532 | -21.1693
13.0265 | -34.7537 | -21.1775
13.0326 | -34.7542 | -21.1858
13.0387 | -34.7546 | -21.1942
13.0448 | -34.7551 | -21.2025
13.0509 | -34.7556 | -21.2107
13.057 | -34.756 | -21.2188
13.0631 | -34.7565 | -21.227
13.0692 | -34.757 | -21.2351
13.0753 | -34.7574 | -21.2434
13.0814 | -34.7579 | -21.2519
13.0875 | -34.7584 | -21.2599
13.0936 | -34.7588 | -21.2683
13.0997 | -34.7593 | -21.2767
13.1058 | -34.7597 | -21.2849
13.112 | -34.7602 | -21.2932
13.1181 | -34.7607 | -21.3014
13.1242 | -34.7611 | -21.3097
13.1303 | -34.7616 | -21.3181
13.1364 | -34.762 | -21.3263
13.1425 | -34.7624 | -21.3343
13.1486 | -34.7629 | -21.3425
13.1547 | -34.7634 | -21.3506
13.1608 | -34.7638 | -21.359
13.1669 | -34.7643 | -21.3672
13.173 | -34.7648 | -21.3755
13.1791 | -34.7652 | -21.3835
13.1852 | -34.7657 | -21.3915
13.1913 | -34.7661 | -21.3996
13.1974 | -34.7666 | -21.408
13.2035 | -34.767 | -21.4162
13.2096 | -34.7675 | -21.4246
13.2157 | -34.768 | -21.433
13.2218 | -34.7684 | -21.4412
13.2279 | -34.7689 | -21.4494
13.234 | -34.7694 | -21.4573
13.2401 | -34.7699 | -21.4656
13.2462 | -34.7703 | -21.4738
13.2524 | -34.7708 | -21.4818
13.2585 | -34.7713 | -21.4897
13.2646 | -34.7717 | -21.4978
13.2707 | -34.7722 | -21.5061
13.2768 | -34.7727 | -21.5145
13.2829 | -34.7732 | -21.5229
13.289 | -34.7736 | -21.531
13.2951 | -34.7741 | -21.5386
13.3012 | -34.7746 | -21.5466
13.3073 | -34.7751 | -21.5551
13.3134 | -34.7755 | -21.5634
13.3195 | -34.7759 | -21.5716
13.3256 | -34.7764 | -21.5799
13.3317 | -34.7769 | -21.5883
13.3378 | -34.7773 | -21.5963
13.3439 | -34.7777 | -21.6048
13.35 | -34.7782 | -21.6132
13.3561 | -34.7786 | -21.6214
13.3622 | -34.779 | -21.6291
13.3683 | -34.7794 | -21.6371
13.3744 | -34.7799 | -21.6452
13.3805 | -34.7803 | -21.6537
13.3866 | -34.7808 | -21.6618
13.3927 | -34.7813 | -21.6699
13.3989 | -34.7817 | -21.6779
13.405 | -34.7821 | -21.6861
13.4111 | -34.7826 | -21.6944
13.4172 | -34.7831 | -21.7024
13.4233 | -34.7836 | -21.711
13.4294 | -34.7841 | -21.7193
13.4355 | -34.7845 | -21.7275
13.4416 | -34.785 | -21.7356
13.4477 | -34.7854 | -21.744
13.4538 | -34.7859 | -21.7521
13.4599 | -34.7864 | -21.7602
13.466 | -34.7868 | -21.7685
13.4721 | -34.7873 | -21.7765
13.4782 | -34.7877 | -21.7848
13.4843 | -34.7882 | -21.7931
13.4904 | -34.7887 | -21.8012
13.4965 | -34.7891 | -21.8095
13.5026 | -34.7896 | -21.818
13.5087 | -34.7901 | -21.8265
13.5148 | -34.7905 | -21.8349
13.5209 | -34.791 | -21.8431
13.527 | -34.7914 | -21.8516
13.5331 | -34.7919 | -21.8598
13.5393 | -34.7923 | -21.8679
13.5454 | -34.7928 | -21.8758
13.5515 | -34.7933 | -21.8843
13.5576 | -34.7937 | -21.8925
13.5637 | -34.7941 | -21.9009
13.5698 | -34.7946 | -21.909
13.5759 | -34.7951 | -21.9171
13.582 | -34.7956 | -21.9253
13.5881 | -34.796 | -21.9334
13.5942 | -34.7965 | -21.9419
13.6003 | -34.7969 | -21.95
13.6064 | -34.7974 | -21.9583
13.6125 | -34.7978 | -21.9664
13.6186 | -34.7983 | -21.9744
13.6247 | -34.7987 | -21.9827
13.6308 | -34.7992 | -21.9909
13.6369 | -34.7997 | -21.9993
13.643 | -34.8001 | -22.0078
13.6491 | -34.8006 | -22.0158
13.6552 | -34.801 | -22.0241
13.6613 | -34.8015 | -22.0324
13.6674 | -34.8019 | -22.0409
13.6735 | -34.8024 | -22.0491
13.6796 | -34.8028 | -22.0572
13.6858 | -34.8033 | -22.0654
13.6919 | -34.8038 | -22.0737
13.698 | -34.8042 | -22.0821
13.7041 | -34.8047 | -22.0904
13.7102 | -34.8051 | -22.0987
13.7163 | -34.8056 | -22.1067
13.7224 | -34.806 | -22.1148
13.7285 | -34.8065 | -22.1232
13.7346 | -34.8069 | -22.1313
13.7407 | -34.8074 | -22.1395
13.7468 | -34.8079 | -22.1477
13.7529 | -34.8083 | -22.1557
13.759 | -34.8087 | -22.1639
13.7651 | -34.8092 | -22.1721
13.7712 | -34.8097 | -22.1803
13.7773 | -34.8101 | -22.1884
13.7834 | -34.8105 | -22.1966
13.7895 | -34.811 | -22.2046
13.7956 | -34.8115 | -22.213
13.8017 | -34.8119 | -22.221
13.8078 | -34.8123 | -22.2291
13.8139 | -34.8128 | -22.2373
13.82 | -34.8132 | -22.2455
13.8262 | -34.8137 | -22.2537
13.8323 | -34.8142 | -22.2621
13.8384 | -34.8146 | -22.2704
13.8445 | -34.8151 | -22.2785
13.8506 | -34.8155 | -22.2865
13.8567 | -34.816 | -22.2949
13.8628 | -34.8164 | -22.3031
13.8689 | -34.8169 | -22.3112
13.875 | -34.8174 | -22.3194
13.8811 | -34.8178 | -22.3276
13.8872 | -34.8182 | -22.3357
13.8933 | -34.8187 | -22.3438
13.8994 | -34.8191 | -22.3519
13.9055 | -34.8196 | -22.36
13.9116 | -34.82 | -22.368
13.9177 | -34.8205 | -22.3761
13.9238 | -34.8209 | -22.3841
13.9299 | -34.8214 | -22.3922
13.936 | -34.8218 | -22.4004
13.9421 | -34.8222 | -22.4085
13.9482 | -34.8227 | -22.4166
13.9543 | -34.8231 | -22.4249
13.9604 | -34.8236 | -22.4333
13.9665 | -34.8241 | -22.4414
13.9727 | -34.8245 | -22.4498
13.9788 | -34.825 | -22.458
13.9849 | -34.8254 | -22.4663
13.991 | -34.8259 | -22.4747
13.9971 | -34.8263 | -22.4828
14.0032 | -34.8268 | -22.4911
14.0093 | -34.8272 | -22.4991
14.0154 | -34.8277 | -22.5072
14.0215 | -34.8281 | -22.5153
14.0276 | -34.8286 | -22.5234
14.0337 | -34.8291 | -22.5316
14.0398 | -34.8295 | -22.5398
14.0459 | -34.83 | -22.5478
14.052 | -34.8304 | -22.556
14.0581 | -34.8309 | -22.5641
14.0642 | -34.8313 | -22.5724
14.0703 | -34.8318 | -22.5808
14.0764 | -34.8323 | -22.5888
14.0825 | -34.8327 | -22.5971
14.0886 | -34.8332 | -22.6054
14.0947 | -34.8336 | -22.6136
14.1008 | -34.8341 | -22.6216
14.1069 | -34.8346 | -22.6294
14.1131 | -34.835 | -22.6378
14.1192 | -34.8354 | -22.646
14.1253 | -34.8359 | -22.654
14.1314 | -34.8364 | -22.6621
14.1375 | -34.8369 | -22.6704
14.1436 | -34.8374 | -22.6786
14.1497 | -34.8378 | -22.6868
14.1558 | -34.8383 | -22.6951
14.1619 | -34.8388 | -22.7033
14.168 | -34.8392 | -22.7117
14.1741 | -34.8397 | -22.7199
14.1802 | -34.8401 | -22.7279
14.1863 | -34.8406 | -22.736
14.1924 | -34.8411 | -22.7442
14.1985 | -34.8416 | -22.7522
14.2046 | -34.842 | -22.7606
14.2107 | -34.8425 | -22.7686
14.2168 | -34.843 | -22.7768
14.2229 | -34.8434 | -22.7848
14.229 | -34.8439 | -22.7933
14.2351 | -34.8444 | -22.8013
14.2412 | -34.8448 | -22.8098
14.2473 | -34.8453 | -22.8179
14.2534 | -34.8457 | -22.8258
14.2596 | -34.8462 | -22.834
14.2657 | -34.8467 | -22.8423
14.2718 | -34.8471 | -22.8504
14.2779 | -34.8476 | -22.8586
14.284 | -34.8481 | -22.8667
14.2901 | -34.8485 | -22.8748
14.2962 | -34.8489 | -22.8831
14.3023 | -34.8494 | -22.8914
14.3084 | -34.8498 | -22.8997
14.3145 | -34.8503 | -22.9077
14.3206 | -34.8508 | -22.9162
14.3267 | -34.8512 | -22.9243
14.3328 | -34.8517 | -22.9324
14.3389 | -34.8522 | -22.9406
14.345 | -34.8526 | -22.9486
14.3511 | -34.8531 | -22.9568
14.3572 | -34.8535 | -22.965
14.3633 | -34.854 | -22.9733
14.3694 | -34.8545 | -22.9812
14.3755 | -34.855 | -22.9895
14.3816 | -34.8554 | -22.9979
14.3877 | -34.8559 | -23.0059
14.3938 | -34.8564 | -23.0144
14.4 | -34.8568 | -23.0224
14.4061 | -34.8573 | -23.0305
14.4122 | -34.8577 | -23.0387
14.4183 | -34.8582 | -23.0467
14.4244 | -34.8587 | -23.0548
14.4305 | -34.8592 | -23.063
14.4366 | -34.8597 | -23.0711
14.4427 | -34.8601 | -23.0795
14.4488 | -34.8606 | -23.0878
14.4549 | -34.8611 | -23.096
14.461 | -34.8615 | -23.1041
14.4671 | -34.862 | -23.1123
14.4732 | -34.8625 | -23.1206
14.4793 | -34.8629 | -23.1287
14.4854 | -34.8634 | -23.1368
14.4915 | -34.8639 | -23.1447
14.4976 | -34.8644 | -23.1528
14.5037 | -34.8649 | -23.161
14.5098 | -34.8653 | -23.1692
14.5159 | -34.8657 | -23.1775
14.522 | -34.8662 | -23.1856
14.5281 | -34.8667 | -23.1936
14.5342 | -34.8672 | -23.2019
14.5403 | -34.8676 | -23.2101
14.5465 | -34.8681 | -23.2183
14.5526 | -34.8685 | -23.2267
14.5587 | -34.869 | -23.235
14.5648 | -34.8695 | -23.243
14.5709 | -34.87 | -23.2511
14.577 | -34.8704 | -23.2592
14.5831 | -34.8709 | -23.2673
14.5892 | -34.8714 | -23.2754
14.5953 | -34.8718 | -23.2834
14.6014 | -34.8723 | -23.2915
14.6075 | -34.8727 | -23.2996
14.6136 | -34.8732 | -23.3078
14.6197 | -34.8736 | -23.3159
14.6258 | -34.8741 | -23.3241
14.6319 | -34.8746 | -23.3321
14.638 | -34.875 | -23.3403
14.6441 | -34.8755 | -23.3485
14.6502 | -34.8759 | -23.3567
14.6563 | -34.8764 | -23.3649
14.6624 | -34.8769 | -23.373
14.6685 | -34.8774 | -23.3809
14.6746 | -34.8778 | -23.3891
14.6807 | -34.8783 | -23.3972
14.6869 | -34.8788 | -23.4052
14.693 | -34.8793 | -23.4133
14.6991 | -34.8797 | -23.4214
14.7052 | -34.8802 | -23.4294
14.7113 | -34.8807 | -23.4377
14.7174 | -34.8811 | -23.4459
14.7235 | -34.8816 | -23.4539
14.7296 | -34.8821 | -23.4619
14.7357 | -34.8826 | -23.47
14.7418 | -34.883 | -23.4782
14.7479 | -34.8835 | -23.4862
14.754 | -34.884 | -23.4942
14.7601 | -34.8845 | -23.5023
14.7662 | -34.8849 | -23.5104
14.7723 | -34.8854 | -23.5183
14.7784 | -34.8859 | -23.5265
14.7845 | -34.8863 | -23.5346
14.7906 | -34.8868 | -23.5429
14.7967 | -34.8872 | -23.5509
14.8028 | -34.8877 | -23.559
14.8089 | -34.8882 | -23.5672
14.815 | -34.8887 | -23.5753
14.8211 | -34.8891 | -23.5834
14.8272 | -34.8896 | -23.5914
14.8334 | -34.8901 | -23.5995
14.8395 | -34.8905 | -23.6075
14.8456 | -34.891 | -23.6156
14.8517 | -34.8915 | -23.6239
14.8578 | -34.892 | -23.6321
14.8639 | -34.8924 | -23.6402
14.87 | -34.8929 | -23.6482
14.8761 | -34.8934 | -23.6562
14.8822 | -34.8938 | -23.6645
14.8883 | -34.8943 | -23.6727
14.8944 | -34.8948 | -23.6806
14.9005 | -34.8952 | -23.6887
14.9066 | -34.8957 | -23.6968
14.9127 | -34.8962 | -23.7048
14.9188 | -34.8966 | -23.7129
14.9249 | -34.8971 | -23.7209
14.931 | -34.8976 | -23.729
14.9371 | -34.898 | -23.7372
14.9432 | -34.8985 | -23.7453
14.9493 | -34.899 | -23.7535
14.9554 | -34.8995 | -23.7617
14.9615 | -34.8999 | -23.7697
14.9676 | -34.9004 | -23.778
14.9738 | -34.9009 | -23.7861
14.9799 | -34.9013 | -23.7944
14.986 | -34.9018 | -23.8024
14.9921 | -34.9022 | -23.8104
14.9982 | -34.9027 | -23.8185
15.0043 | -34.9032 | -23.8265
15.0104 | -34.9036 | -23.8348
15.0165 | -34.9041 | -23.843
15.0226 | -34.9046 | -23.851
15.0287 | -34.9051 | -23.8591
15.0348 | -34.9056 | -23.8673
15.0409 | -34.9061 | -23.8753
15.047 | -34.9065 | -23.8834
15.0531 | -34.907 | -23.8914
15.0592 | -34.9075 | -23.8996
15.0653 | -34.908 | -23.9077
15.0714 | -34.9085 | -23.9157
15.0775 | -34.9089 | -23.9236
15.0836 | -34.9094 | -23.9317
15.0897 | -34.9099 | -23.9397
15.0958 | -34.9104 | -23.9477
15.1019 | -34.9109 | -23.9556
15.108 | -34.9114 | -23.9636
15.1141 | -34.9118 | -23.9716
15.1203 | -34.9123 | -23.9798
15.1264 | -34.9128 | -23.9877
15.1325 | -34.9133 | -23.9959
15.1386 | -34.9137 | -24.0039
15.1447 | -34.9142 | -24.0119
15.1508 | -34.9147 | -24.0199
15.1569 | -34.9152 | -24.0281
15.163 | -34.9157 | -24.0359
15.1691 | -34.9162 | -24.044
15.1752 | -34.9167 | -24.052
15.1813 | -34.9172 | -24.0601
15.1874 | -34.9176 | -24.0681
15.1935 | -34.9181 | -24.076
15.1996 | -34.9186 | -24.0839
15.2057 | -34.9191 | -24.092
15.2118 | -34.9196 | -24.1001
15.2179 | -34.9201 | -24.1082
15.224 | -34.9205 | -24.1161
15.2301 | -34.921 | -24.1241
15.2362 | -34.9215 | -24.1321
15.2423 | -34.922 | -24.1403
15.2484 | -34.9225 | -24.1484
15.2545 | -34.923 | -24.1566
15.2607 | -34.9235 | -24.1646
15.2668 | -34.924 | -24.1725
15.2729 | -34.9245 | -24.1805
15.279 | -34.925 | -24.1887
15.2851 | -34.9255 | -24.1966
15.2912 | -34.9259 | -24.2047
15.2973 | -34.9264 | -24.2128
15.3034 | -34.9269 | -24.2208
15.3095 | -34.9274 | -24.2288
15.3156 | -34.9279 | -24.2366
15.3217 | -34.9283 | -24.2447
15.3278 | -34.9288 | -24.2529
15.3339 | -34.9293 | -24.2607
15.34 | -34.9298 | -24.2685
15.3461 | -34.9303 | -24.2766
15.3522 | -34.9308 | -24.2846
15.3583 | -34.9312 | -24.2928
15.3644 | -34.9317 | -24.3006
15.3705 | -34.9322 | -24.3084
15.3766 | -34.9327 | -24.3163
15.3827 | -34.9332 | -24.3243
15.3888 | -34.9337 | -24.3323
15.3949 | -34.9342 | -24.3404
15.401 | -34.9347 | -24.3484
15.4072 | -34.9351 | -24.3562
15.4133 | -34.9356 | -24.3642
15.4194 | -34.9361 | -24.3723
15.4255 | -34.9366 | -24.3803
15.4316 | -34.9371 | -24.3884
15.4377 | -34.9376 | -24.3964
15.4438 | -34.9381 | -24.4044
15.4499 | -34.9386 | -24.4122
15.456 | -34.939 | -24.4201
15.4621 | -34.9395 | -24.4279
15.4682 | -34.94 | -24.4361
15.4743 | -34.9405 | -24.4438
15.4804 | -34.941 | -24.4519
15.4865 | -34.9414 | -24.4598
15.4926 | -34.9419 | -24.4677
15.4987 | -34.9424 | -24.4755
15.5048 | -34.9429 | -24.4835
15.5109 | -34.9434 | -24.4915
15.517 | -34.9439 | -24.4995
15.5231 | -34.9444 | -24.5075
15.5292 | -34.9449 | -24.5155
15.5353 | -34.9453 | -24.5236
15.5414 | -34.9458 | -24.5315
15.5476 | -34.9464 | -24.5392
15.5537 | -34.9469 | -24.5473
15.5598 | -34.9473 | -24.5555
15.5659 | -34.9478 | -24.5635
15.572 | -34.9483 | -24.5713
15.5781 | -34.9488 | -24.5791
15.5842 | -34.9493 | -24.5869
15.5903 | -34.9498 | -24.595
15.5964 | -34.9502 | -24.6029
15.6025 | -34.9507 | -24.6108
15.6086 | -34.9512 | -24.6188
15.6147 | -34.9517 | -24.6268
15.6208 | -34.9522 | -24.6344
15.6269 | -34.9527 | -24.6424
15.633 | -34.9532 | -24.6502
15.6391 | -34.9537 | -24.6582
15.6452 | -34.9541 | -24.6662
15.6513 | -34.9547 | -24.6739
15.6574 | -34.9552 | -24.6819
15.6635 | -34.9557 | -24.6899
15.6696 | -34.9561 | -24.6979
15.6757 | -34.9566 | -24.7058
15.6818 | -34.9571 | -24.7138
15.688 | -34.9576 | -24.7217
15.6941 | -34.9581 | -24.7296
15.7002 | -34.9586 | -24.7375
15.7063 | -34.959 | -24.7454
15.7124 | -34.9595 | -24.7535
15.7185 | -34.96 | -24.7614
15.7246 | -34.9605 | -24.7693
15.7307 | -34.961 | -24.7772
15.7368 | -34.9615 | -24.7852
15.7429 | -34.9619 | -24.793
15.749 | -34.9624 | -24.8009
15.7551 | -34.9629 | -24.8088
15.7612 | -34.9634 | -24.8165
15.7673 | -34.9639 | -24.8243
15.7734 | -34.9644 | -24.832
15.7795 | -34.9649 | -24.84
15.7856 | -34.9653 | -24.8479
15.7917 | -34.9658 | -24.8554
15.7978 | -34.9663 | -24.8634
15.8039 | -34.9668 | -24.8712
15.81 | -34.9673 | -24.879
15.8161 | -34.9678 | -24.8869
15.8222 | -34.9683 | -24.8948
15.8283 | -34.9688 | -24.9023
15.8345 | -34.9693 | -24.9101
15.8406 | -34.9698 | -24.9179
15.8467 | -34.9703 | -24.9256
15.8528 | -34.9707 | -24.9336
15.8589 | -34.9712 | -24.9414
15.865 | -34.9717 | -24.9489
15.8711 | -34.9722 | -24.957
15.8772 | -34.9727 | -24.9647
15.8833 | -34.9732 | -24.9725
15.8894 | -34.9737 | -24.9804
15.8955 | -34.9742 | -24.9883
15.9016 | -34.9747 | -24.9961
15.9077 | -34.9751 | -25.0038
15.9138 | -34.9756 | -25.0118
15.9199 | -34.9761 | -25.0196
15.926 | -34.9766 | -25.0276
15.9321 | -34.9771 | -25.0353
15.9382 | -34.9776 | -25.0429
15.9443 | -34.9781 | -25.0507
15.9504 | -34.9785 | -25.0586
15.9565 | -34.979 | -25.0663
15.9626 | -34.9795 | -25.0741
15.9687 | -34.98 | -25.0816
15.9749 | -34.9804 | -25.0892
15.981 | -34.9809 | -25.0973
15.9871 | -34.9814 | -25.1051
15.9932 | -34.9818 | -25.1129
15.9993 | -34.9823 | -25.1206
16.0054 | -34.9828 | -25.1282
16.0115 | -34.9833 | -25.1361
16.0176 | -34.9838 | -25.144
16.0237 | -34.9843 | -25.1519
16.0298 | -34.9848 | -25.1597
16.0359 | -34.9853 | -25.1674
16.042 | -34.9858 | -25.1751
16.0481 | -34.9862 | -25.1828
16.0542 | -34.9867 | -25.1906
16.0603 | -34.9872 | -25.1983
16.0664 | -34.9877 | -25.2062
16.0725 | -34.9882 | -25.2138
16.0786 | -34.9887 | -25.2216
16.0847 | -34.9892 | -25.2294
16.0908 | -34.9897 | -25.2372
16.0969 | -34.9901 | -25.2451
16.103 | -34.9906 | -25.253
16.1091 | -34.9912 | -25.2605
16.1152 | -34.9916 | -25.2682
16.1214 | -34.9921 | -25.2761
16.1275 | -34.9926 | -25.2839
16.1336 | -34.9931 | -25.2914
16.1397 | -34.9936 | -25.2992
16.1458 | -34.9941 | -25.3067
16.1519 | -34.9946 | -25.3146
16.158 | -34.995 | -25.3222
16.1641 | -34.9955 | -25.33
16.1702 | -34.996 | -25.3376
16.1763 | -34.9965 | -25.3455
16.1824 | -34.997 | -25.353
16.1885 | -34.9974 | -25.3607
16.1946 | -34.998 | -25.3685
16.2007 | -34.9985 | -25.3762
16.2068 | -34.9989 | -25.3839
16.2129 | -34.9994 | -25.3917
16.219 | -34.9999 | -25.3991
16.2251 | -35.0004 | -25.4067
16.2312 | -35.0009 | -25.4144
16.2373 | -35.0013 | -25.4221
16.2434 | -35.0018 | -25.4298
16.2495 | -35.0023 | -25.4375
16.2556 | -35.0028 | -25.4452
16.2618 | -35.0033 | -25.4528
16.2679 | -35.0038 | -25.4606
16.274 | -35.0043 | -25.4684
16.2801 | -35.0048 | -25.4761
16.2862 | -35.0053 | -25.4835
16.2923 | -35.0058 | -25.4912
16.2984 | -35.0062 | -25.499
16.3045 | -35.0068 | -25.5066
16.3106 | -35.0072 | -25.5143
16.3167 | -35.0077 | -25.5221
16.3228 | -35.0082 | -25.5297
16.3289 | -35.0086 | -25.5372
16.335 | -35.0091 | -25.5449
16.3411 | -35.0096 | -25.5528
16.3472 | -35.0101 | -25.5605
16.3533 | -35.0106 | -25.5681
16.3594 | -35.0111 | -25.5757
16.3655 | -35.0116 | -25.5835
16.3716 | -35.0121 | -25.5909
16.3777 | -35.0126 | -25.5987
16.3838 | -35.013 | -25.6064
16.3899 | -35.0135 | -25.6138
16.396 | -35.014 | -25.6215
16.4021 | -35.0145 | -25.6291
16.4083 | -35.015 | -25.6367
16.4144 | -35.0155 | -25.6446
16.4205 | -35.016 | -25.6522
16.4266 | -35.0164 | -25.6596
16.4327 | -35.0169 | -25.6672
16.4388 | -35.0174 | -25.6748
16.4449 | -35.0179 | -25.6827
16.451 | -35.0183 | -25.6903
16.4571 | -35.0188 | -25.6981
16.4632 | -35.0193 | -25.7055
16.4693 | -35.0198 | -25.7132
16.4754 | -35.0203 | -25.7211
16.4815 | -35.0208 | -25.7286
16.4876 | -35.0213 | -25.736
16.4937 | -35.0218 | -25.7437
16.4998 | -35.0223 | -25.751
16.5059 | -35.0228 | -25.7585
16.512 | -35.0233 | -25.7664
16.5181 | -35.0238 | -25.7741
16.5242 | -35.0243 | -25.7819
16.5303 | -35.0247 | -25.7894
16.5364 | -35.0252 | -25.7971
16.5425 | -35.0257 | -25.8046
16.5487 | -35.0261 | -25.8121
16.5548 | -35.0266 | -25.8199
16.5609 | -35.0271 | -25.8277
16.567 | -35.0276 | -25.8353
16.5731 | -35.0281 | -25.843
16.5792 | -35.0286 | -25.8506
16.5853 | -35.029 | -25.8583
16.5914 | -35.0295 | -25.8658
16.5975 | -35.03 | -25.8735
16.6036 | -35.0305 | -25.8811
16.6097 | -35.0309 | -25.8886
16.6158 | -35.0314 | -25.8962
16.6219 | -35.0319 | -25.9038
16.628 | -35.0323 | -25.9115
16.6341 | -35.0328 | -25.919
16.6402 | -35.0333 | -25.9263
16.6463 | -35.0337 | -25.9339
16.6524 | -35.0342 | -25.9414
16.6585 | -35.0347 | -25.9491
16.6646 | -35.0351 | -25.9566
16.6707 | -35.0356 | -25.9642
16.6768 | -35.0361 | -25.9719
16.6829 | -35.0365 | -25.9793
16.689 | -35.037 | -25.987
16.6952 | -35.0375 | -25.9945
16.7013 | -35.038 | -26.0021
16.7074 | -35.0384 | -26.0094
16.7135 | -35.0389 | -26.0169
16.7196 | -35.0394 | -26.0245
16.7257 | -35.0399 | -26.0321
16.7318 | -35.0404 | -26.0397
16.7379 | -35.0409 | -26.0472
16.744 | -35.0413 | -26.0544
16.7501 | -35.0418 | -26.062
16.7562 | -35.0423 | -26.0695
16.7623 | -35.0428 | -26.0769
16.7684 | -35.0433 | -26.0843
16.7745 | -35.0437 | -26.0918
16.7806 | -35.0442 | -26.0989
16.7867 | -35.0447 | -26.1062
16.7928 | -35.0452 | -26.1136
16.7989 | -35.0457 | -26.1211
16.805 | -35.0462 | -26.1289
16.8111 | -35.0467 | -26.1363
16.8172 | -35.0472 | -26.1438
16.8233 | -35.0477 | -26.1512
16.8294 | -35.0481 | -26.1588
16.8356 | -35.0486 | -26.1664
16.8417 | -35.0491 | -26.1741
16.8478 | -35.0496 | -26.1815
16.8539 | -35.0501 | -26.189
16.86 | -35.0506 | -26.1968
16.8661 | -35.0511 | -26.2044
16.8722 | -35.0515 | -26.2119
16.8783 | -35.052 | -26.2194
16.8844 | -35.0525 | -26.2267
16.8905 | -35.053 | -26.2342
16.8966 | -35.0535 | -26.2417
16.9027 | -35.054 | -26.2492
16.9088 | -35.0545 | -26.2568
16.9149 | -35.0549 | -26.2644
16.921 | -35.0554 | -26.2714
16.9271 | -35.0559 | -26.279
16.9332 | -35.0564 | -26.2866
16.9393 | -35.0569 | -26.2943
16.9454 | -35.0573 | -26.3018
16.9515 | -35.0578 | -26.3093
16.9576 | -35.0583 | -26.3164
16.9637 | -35.0588 | -26.3237
16.9698 | -35.0593 | -26.3313
16.9759 | -35.0598 | -26.3389
16.9821 | -35.0603 | -26.3464
16.9882 | -35.0608 | -26.3538
16.9943 | -35.0613 | -26.3613
17.0004 | -35.0617 | -26.3687
17.0065 | -35.0622 | -26.3762
17.0126 | -35.0627 | -26.3837
17.0187 | -35.0632 | -26.3913
17.0248 | -35.0636 | -26.3985
17.0309 | -35.0641 | -26.4059
17.037 | -35.0646 | -26.4131
17.0431 | -35.065 | -26.4206
17.0492 | -35.0655 | -26.428
17.0553 | -35.066 | -26.4354
17.0614 | -35.0665 | -26.4429
17.0675 | -35.067 | -26.4502
17.0736 | -35.0674 | -26.4577
17.0797 | -35.0679 | -26.465
17.0858 | -35.0684 | -26.4724
17.0919 | -35.0689 | -26.4799
17.098 | -35.0693 | -26.4871
17.1041 | -35.0698 | -26.4945
17.1102 | -35.0703 | -26.5022
17.1163 | -35.0708 | -26.5098
17.1225 | -35.0713 | -26.5174
17.1286 | -35.0718 | -26.5246
17.1347 | -35.0723 | -26.532
17.1408 | -35.0728 | -26.5396
17.1469 | -35.0732 | -26.547
17.153 | -35.0737 | -26.5542
17.1591 | -35.0742 | -26.5616
17.1652 | -35.0747 | -26.5689
17.1713 | -35.0752 | -26.5762
17.1774 | -35.0756 | -26.5836
17.1835 | -35.0761 | -26.591
17.1896 | -35.0766 | -26.5983
17.1957 | -35.077 | -26.6059
17.2018 | -35.0775 | -26.6133
17.2079 | -35.078 | -26.6209
17.214 | -35.0784 | -26.6285
17.2201 | -35.0789 | -26.6359
17.2262 | -35.0794 | -26.6434
17.2323 | -35.0799 | -26.651
17.2384 | -35.0804 | -26.6583
17.2445 | -35.0808 | -26.6656
17.2506 | -35.0813 | -26.673
17.2567 | -35.0818 | -26.6805
17.2628 | -35.0823 | -26.6879
17.269 | -35.0828 | -26.6953
17.2751 | -35.0832 | -26.7028
17.2812 | -35.0837 | -26.71
17.2873 | -35.0842 | -26.7174
17.2934 | -35.0847 | -26.7249
17.2995 | -35.0851 | -26.7323
17.3056 | -35.0856 | -26.7397
17.3117 | -35.0861 | -26.7469
17.3178 | -35.0865 | -26.7542
17.3239 | -35.087 | -26.7617
17.33 | -35.0874 | -26.7691
17.3361 | -35.0879 | -26.7765
17.3422 | -35.0884 | -26.7837
17.3483 | -35.0888 | -26.7913
17.3544 | -35.0893 | -26.7987
17.3605 | -35.0898 | -26.8062
17.3666 | -35.0902 | -26.8136
17.3727 | -35.0907 | -26.821
17.3788 | -35.0911 | -26.8279
17.3849 | -35.0916 | -26.8351
17.391 | -35.0921 | -26.8425
17.3971 | -35.0926 | -26.8498
17.4032 | -35.0931 | -26.8572
17.4094 | -35.0936 | -26.8645
17.4155 | -35.094 | -26.8718
17.4216 | -35.0945 | -26.879
17.4277 | -35.095 | -26.8864
17.4338 | -35.0955 | -26.8938
17.4399 | -35.096 | -26.9012
17.446 | -35.0964 | -26.9082
17.4521 | -35.0969 | -26.9156
17.4582 | -35.0974 | -26.9227
17.4643 | -35.0978 | -26.93
17.4704 | -35.0983 | -26.9372
17.4765 | -35.0987 | -26.9446
17.4826 | -35.0992 | -26.9518
17.4887 | -35.0997 | -26.9593
17.4948 | -35.1001 | -26.9666
17.5009 | -35.1006 | -26.9738
17.507 | -35.1011 | -26.9812
17.5131 | -35.1015 | -26.9886
17.5192 | -35.102 | -26.9959
17.5253 | -35.1024 | -27.003
17.5314 | -35.1028 | -27.0104
17.5375 | -35.1033 | -27.0179
17.5436 | -35.1038 | -27.0254
17.5497 | -35.1043 | -27.0325
17.5559 | -35.1047 | -27.0396
17.562 | -35.1052 | -27.0468
17.5681 | -35.1057 | -27.054
17.5742 | -35.1061 | -27.0615
17.5803 | -35.1066 | -27.0689
17.5864 | -35.1071 | -27.0761
17.5925 | -35.1076 | -27.0834
17.5986 | -35.108 | -27.0906
17.6047 | -35.1085 | -27.0982
17.6108 | -35.109 | -27.1054
17.6169 | -35.1095 | -27.1127
17.623 | -35.1099 | -27.1199
17.6291 | -35.1104 | -27.1272
17.6352 | -35.1108 | -27.1348
17.6413 | -35.1113 | -27.142
17.6474 | -35.1118 | -27.1494
17.6535 | -35.1122 | -27.1566
17.6596 | -35.1127 | -27.1639
17.6657 | -35.1132 | -27.1711
17.6718 | -35.1136 | -27.1784
17.6779 | -35.1141 | -27.1854
17.684 | -35.1146 | -27.1928
17.6901 | -35.115 | -27.2
17.6963 | -35.1155 | -27.2069
17.7024 | -35.1159 | -27.2141
17.7085 | -35.1164 | -27.2211
17.7146 | -35.1169 | -27.2283
17.7207 | -35.1173 | -27.2356
17.7268 | -35.1178 | -27.2425
17.7329 | -35.1183 | -27.2498
17.739 | -35.1188 | -27.2572
17.7451 | -35.1192 | -27.2645
17.7512 | -35.1197 | -27.2716
17.7573 | -35.1202 | -27.2789
17.7634 | -35.1206 | -27.286
17.7695 | -35.1211 | -27.2933
17.7756 | -35.1216 | -27.3006
17.7817 | -35.122 | -27.3078
17.7878 | -35.1225 | -27.3153
17.7939 | -35.123 | -27.3226
17.8 | -35.1235 | -27.3298
17.8061 | -35.124 | -27.3374
17.8122 | -35.1245 | -27.3445
17.8183 | -35.1249 | -27.3518
17.8244 | -35.1254 | -27.3592
17.8305 | -35.1258 | -27.3666
17.8366 | -35.1263 | -27.3732
17.8428 | -35.1267 | -27.3807
17.8489 | -35.1272 | -27.3879
17.855 | -35.1277 | -27.3953
17.8611 | -35.1282 | -27.4025
17.8672 | -35.1286 | -27.4096
17.8733 | -35.1291 | -27.4169
17.8794 | -35.1296 | -27.4241
17.8855 | -35.13 | -27.4314
17.8916 | -35.1305 | -27.4386
17.8977 | -35.1309 | -27.4459
17.9038 | -35.1314 | -27.4529
17.9099 | -35.1318 | -27.4601
17.916 | -35.1323 | -27.4674
17.9221 | -35.1328 | -27.4747
17.9282 | -35.1333 | -27.4818
17.9343 | -35.1337 | -27.4893
17.9404 | -35.1342 | -27.4965
17.9465 | -35.1346 | -27.5037
17.9526 | -35.1351 | -27.5107
17.9587 | -35.1356 | -27.518
17.9648 | -35.136 | -27.5251
17.9709 | -35.1365 | -27.5325
17.977 | -35.137 | -27.5396
17.9832 | -35.1375 | -27.547
17.9893 | -35.1379 | -27.5543
17.9954 | -35.1384 | -27.5617
18.0015 | -35.1388 | -27.569
18.0076 | -35.1393 | -27.5761
18.0137 | -35.1397 | -27.5831
18.0198 | -35.1402 | -27.5901
18.0259 | -35.1407 | -27.5973
18.032 | -35.1411 | -27.6046
18.0381 | -35.1416 | -27.6119
18.0442 | -35.142 | -27.619
18.0503 | -35.1425 | -27.6262
18.0564 | -35.1429 | -27.6333
18.0625 | -35.1434 | -27.6407
18.0686 | -35.1438 | -27.6481
18.0747 | -35.1443 | -27.6552
18.0808 | -35.1447 | -27.6621
18.0869 | -35.1452 | -27.6695
18.093 | -35.1456 | -27.6765
18.0991 | -35.1461 | -27.6838
18.1052 | -35.1465 | -27.6911
18.1113 | -35.147 | -27.6984
18.1174 | -35.1474 | -27.7053
18.1236 | -35.1479 | -27.7123
18.1297 | -35.1484 | -27.7194
18.1358 | -35.1488 | -27.7265
18.1419 | -35.1493 | -27.7338
18.148 | -35.1498 | -27.741
18.1541 | -35.1503 | -27.7479
18.1602 | -35.1508 | -27.7548
18.1663 | -35.1512 | -27.762
18.1724 | -35.1517 | -27.7692
18.1785 | -35.1522 | -27.7761
18.1846 | -35.1526 | -27.7832
18.1907 | -35.1531 | -27.7904
18.1968 | -35.1535 | -27.7976
18.2029 | -35.154 | -27.805
18.209 | -35.1544 | -27.8122
18.2151 | -35.1549 | -27.8196
18.2212 | -35.1554 | -27.8266
18.2273 | -35.1558 | -27.8339
18.2334 | -35.1563 | -27.8409
18.2395 | -35.1568 | -27.848
18.2456 | -35.1573 | -27.8553
18.2517 | -35.1577 | -27.8625
18.2578 | -35.1582 | -27.8696
18.2639 | -35.1587 | -27.877
18.2701 | -35.1591 | -27.8841
18.2762 | -35.1596 | -27.8911
18.2823 | -35.1601 | -27.8982
18.2884 | -35.1605 | -27.9053
18.2945 | -35.161 | -27.9122
18.3006 | -35.1614 | -27.9195
18.3067 | -35.1619 | -27.9268
18.3128 | -35.1624 | -27.934
18.3189 | -35.1628 | -27.9413
18.325 | -35.1633 | -27.9484
18.3311 | -35.1637 | -27.9556
18.3372 | -35.1641 | -27.9628
18.3433 | -35.1646 | -27.9698
18.3494 | -35.1651 | -27.9766
18.3555 | -35.1655 | -27.9839
18.3616 | -35.166 | -27.9908
18.3677 | -35.1664 | -27.998
18.3738 | -35.1669 | -28.0052
18.3799 | -35.1674 | -28.0122
18.386 | -35.1678 | -28.0193
18.3921 | -35.1682 | -28.0264
18.3982 | -35.1687 | -28.0332
18.4043 | -35.1691 | -28.04
18.4105 | -35.1696 | -28.0474
18.4166 | -35.17 | -28.0544
18.4227 | -35.1705 | -28.0614
18.4288 | -35.1709 | -28.0684
18.4349 | -35.1714 | -28.0753
18.441 | -35.1718 | -28.0823
18.4471 | -35.1723 | -28.0891
18.4532 | -35.1727 | -28.0961
18.4593 | -35.1732 | -28.103
18.4654 | -35.1736 | -28.11
18.4715 | -35.1741 | -28.117
18.4776 | -35.1745 | -28.1243
18.4837 | -35.1749 | -28.1316
18.4898 | -35.1754 | -28.1388
18.4959 | -35.1758 | -28.1461
18.502 | -35.1763 | -28.153
18.5081 | -35.1767 | -28.1602
18.5142 | -35.1771 | -28.1673
18.5203 | -35.1776 | -28.1743
18.5264 | -35.178 | -28.1812
18.5325 | -35.1785 | -28.1885
18.5386 | -35.1789 | -28.1953
18.5447 | -35.1794 | -28.2025
18.5508 | -35.1798 | -28.2094
18.557 | -35.1803 | -28.2166
18.5631 | -35.1807 | -28.2237
18.5692 | -35.1811 | -28.2307
18.5753 | -35.1816 | -28.2374
18.5814 | -35.1821 | -28.2445
18.5875 | -35.1825 | -28.2515
18.5936 | -35.183 | -28.2585
18.5997 | -35.1834 | -28.2655
18.6058 | -35.1839 | -28.2724
18.6119 | -35.1843 | -28.2794
18.618 | -35.1848 | -28.2863
18.6241 | -35.1852 | -28.2936
18.6302 | -35.1857 | -28.3007
18.6363 | -35.1862 | -28.3077
18.6424 | -35.1866 | -28.3144
18.6485 | -35.187 | -28.3214
18.6546 | -35.1875 | -28.3283
18.6607 | -35.1879 | -28.3354
18.6668 | -35.1884 | -28.3427
18.6729 | -35.1889 | -28.35
18.679 | -35.1894 | -28.3568
18.6851 | -35.1898 | -28.3637
18.6912 | -35.1903 | -28.3709
18.6974 | -35.1907 | -28.3778
18.7035 | -35.1912 | -28.3849
18.7096 | -35.1917 | -28.392
18.7157 | -35.1921 | -28.399
18.7218 | -35.1926 | -28.4061
18.7279 | -35.193 | -28.4133
18.734 | -35.1935 | -28.4206
18.7401 | -35.1939 | -28.4276
18.7462 | -35.1944 | -28.4344
18.7523 | -35.1948 | -28.4414
18.7584 | -35.1953 | -28.4483
18.7645 | -35.1958 | -28.4555
18.7706 | -35.1962 | -28.4624
18.7767 | -35.1967 | -28.4693
18.7828 | -35.1971 | -28.4761
18.7889 | -35.1976 | -28.4832
18.795 | -35.198 | -28.4904
18.8011 | -35.1985 | -28.4973
18.8072 | -35.1989 | -28.5046
18.8133 | -35.1994 | -28.5115
18.8194 | -35.1998 | -28.5182
18.8255 | -35.2002 | -28.5252
18.8316 | -35.2006 | -28.5323
18.8377 | -35.2011 | -28.5394
18.8439 | -35.2015 | -28.5462
18.85 | -35.2019 | -28.5534
18.8561 | -35.2024 | -28.5603
18.8622 | -35.2029 | -28.5673
18.8683 | -35.2033 | -28.5742
18.8744 | -35.2037 | -28.5812
18.8805 | -35.2042 | -28.5883
18.8866 | -35.2046 | -28.5956
18.8927 | -35.205 | -28.6024
18.8988 | -35.2055 | -28.6094
18.9049 | -35.206 | -28.6166
18.911 | -35.2064 | -28.6235
18.9171 | -35.2069 | -28.6303
18.9232 | -35.2073 | -28.6369
18.9293 | -35.2078 | -28.6437
18.9354 | -35.2082 | -28.6504
18.9415 | -35.2087 | -28.6571
18.9476 | -35.2091 | -28.6641
18.9537 | -35.2096 | -28.6713
18.9598 | -35.21 | -28.6782
18.9659 | -35.2104 | -28.685
18.972 | -35.2109 | -28.6918
18.9781 | -35.2114 | -28.6987
18.9843 | -35.2118 | -28.7059
18.9904 | -35.2123 | -28.7127
18.9965 | -35.2127 | -28.7192
19.0026 | -35.2132 | -28.7265
19.0087 | -35.2137 | -28.7337
19.0148 | -35.2142 | -28.7406
19.0209 | -35.2146 | -28.7476
19.027 | -35.215 | -28.7546
19.0331 | -35.2155 | -28.7611
19.0392 | -35.2159 | -28.7683
19.0453 | -35.2164 | -28.7753
19.0514 | -35.2168 | -28.7824
19.0575 | -35.2173 | -28.7893
19.0636 | -35.2177 | -28.7958
19.0697 | -35.2181 | -28.8029
19.0758 | -35.2185 | -28.8096
19.0819 | -35.219 | -28.8165
19.088 | -35.2194 | -28.8234
19.0941 | -35.2198 | -28.8304
19.1002 | -35.2203 | -28.8372
19.1063 | -35.2207 | -28.844
19.1124 | -35.2212 | -28.8508
19.1185 | -35.2216 | -28.8579
19.1246 | -35.2221 | -28.865
19.1308 | -35.2225 | -28.8717
19.1369 | -35.223 | -28.8785
19.143 | -35.2235 | -28.8856
19.1491 | -35.2239 | -28.8924
19.1552 | -35.2244 | -28.8996
19.1613 | -35.2248 | -28.9065
19.1674 | -35.2252 | -28.9133
19.1735 | -35.2257 | -28.9202
19.1796 | -35.2261 | -28.9271
19.1857 | -35.2266 | -28.9338
19.1918 | -35.227 | -28.9406
19.1979 | -35.2274 | -28.9475
19.204 | -35.2278 | -28.9543
19.2101 | -35.2283 | -28.9612
19.2162 | -35.2287 | -28.9682
19.2223 | -35.2292 | -28.9751
19.2284 | -35.2296 | -28.9819
19.2345 | -35.23 | -28.9887
19.2406 | -35.2305 | -28.9954
19.2467 | -35.2309 | -29.0023
19.2528 | -35.2314 | -29.0096
19.2589 | -35.2318 | -29.0163
19.265 | -35.2323 | -29.0233
19.2712 | -35.2328 | -29.0301
19.2773 | -35.2333 | -29.0369
19.2834 | -35.2337 | -29.0437
19.2895 | -35.2341 | -29.0508
19.2956 | -35.2345 | -29.0575
19.3017 | -35.235 | -29.0645
19.3078 | -35.2354 | -29.0713
19.3139 | -35.2359 | -29.078
19.32 | -35.2364 | -29.0847
19.3261 | -35.2368 | -29.0917
19.3322 | -35.2372 | -29.0985
19.3383 | -35.2377 | -29.1054
19.3444 | -35.2381 | -29.1121
19.3505 | -35.2385 | -29.119
19.3566 | -35.2389 | -29.1256
19.3627 | -35.2394 | -29.1323
19.3688 | -35.2398 | -29.1392
19.3749 | -35.2403 | -29.1461
19.381 | -35.2407 | -29.1528
19.3871 | -35.2411 | -29.1594
19.3932 | -35.2416 | -29.1664
19.3993 | -35.242 | -29.1736
19.4054 | -35.2425 | -29.1806
19.4115 | -35.2429 | -29.1877
19.4177 | -35.2434 | -29.1941
19.4238 | -35.2438 | -29.2005
19.4299 | -35.2443 | -29.2074
19.436 | -35.2447 | -29.2148
19.4421 | -35.2451 | -29.2215
19.4482 | -35.2456 | -29.2283
19.4543 | -35.2461 | -29.2346
19.4604 | -35.2465 | -29.2414
19.4665 | -35.247 | -29.248
19.4726 | -35.2474 | -29.2547
19.4787 | -35.2479 | -29.2614
19.4848 | -35.2484 | -29.2683
19.4909 | -35.2488 | -29.2751
19.497 | -35.2493 | -29.282
19.5031 | -35.2497 | -29.2888
19.5092 | -35.2502 | -29.2955
19.5153 | -35.2506 | -29.3025
19.5214 | -35.251 | -29.3092
19.5275 | -35.2515 | -29.316
19.5336 | -35.2519 | -29.3224
19.5397 | -35.2523 | -29.3291
19.5458 | -35.2528 | -29.3357
19.5519 | -35.2532 | -29.3425
19.5581 | -35.2536 | -29.3492
19.5642 | -35.2541 | -29.3559
19.5703 | -35.2545 | -29.3626
19.5764 | -35.2549 | -29.3695
19.5825 | -35.2553 | -29.3763
19.5886 | -35.2558 | -29.3833
19.5947 | -35.2562 | -29.3897
19.6008 | -35.2567 | -29.3963
19.6069 | -35.2572 | -29.4032
19.613 | -35.2576 | -29.4101
19.6191 | -35.2581 | -29.4167
19.6252 | -35.2585 | -29.4233
19.6313 | -35.259 | -29.4295
19.6374 | -35.2594 | -29.436
19.6435 | -35.2598 | -29.443
19.6496 | -35.2603 | -29.4498
19.6557 | -35.2607 | -29.4567
19.6618 | -35.2611 | -29.4633
19.6679 | -35.2615 | -29.4704
19.674 | -35.262 | -29.4769
19.6801 | -35.2623 | -29.4837
19.6862 | -35.2628 | -29.4903
19.6923 | -35.2633 | -29.4972
19.6984 | -35.2636 | -29.5036
19.7046 | -35.2641 | -29.5104
19.7107 | -35.2645 | -29.5172
19.7168 | -35.265 | -29.5242
19.7229 | -35.2654 | -29.5308
19.729 | -35.2658 | -29.5378
19.7351 | -35.2663 | -29.5443
19.7412 | -35.2668 | -29.5511
19.7473 | -35.2672 | -29.5578
19.7534 | -35.2677 | -29.5647
19.7595 | -35.2681 | -29.5718
19.7656 | -35.2686 | -29.5788
19.7717 | -35.269 | -29.5847
19.7778 | -35.2694 | -29.5917
19.7839 | -35.2699 | -29.5985
19.79 | -35.2703 | -29.605
19.7961 | -35.2707 | -29.6116
19.8022 | -35.2711 | -29.6178
19.8083 | -35.2716 | -29.6241
19.8144 | -35.2721 | -29.6309
19.8205 | -35.2725 | -29.6374
19.8266 | -35.2729 | -29.644
19.8327 | -35.2733 | -29.6504
19.8388 | -35.2737 | -29.6569
19.845 | -35.2742 | -29.6633
19.8511 | -35.2746 | -29.6701
19.8572 | -35.275 | -29.677
19.8633 | -35.2754 | -29.6836
19.8694 | -35.2758 | -29.6901
19.8755 | -35.2762 | -29.6966
19.8816 | -35.2767 | -29.7031
19.8877 | -35.2771 | -29.7101
19.8938 | -35.2775 | -29.7171
19.8999 | -35.2779 | -29.7238
19.906 | -35.2784 | -29.7306
19.9121 | -35.2788 | -29.7369
19.9182 | -35.2792 | -29.7437
19.9243 | -35.2796 | -29.7504
19.9304 | -35.28 | -29.7571
19.9365 | -35.2805 | -29.7636
19.9426 | -35.2809 | -29.7699
19.9487 | -35.2813 | -29.7766
19.9548 | -35.2818 | -29.7832
19.9609 | -35.2822 | -29.7898
19.967 | -35.2826 | -29.7965
19.9731 | -35.283 | -29.8032
19.9792 | -35.2835 | -29.8097
19.9853 | -35.2839 | -29.8158
19.9915 | -35.2844 | -29.8226
19.9976 | -35.2848 | -29.8295
20.0037 | -35.2852 | -29.8361
20.0098 | -35.2856 | -29.8424
20.0159 | -35.2862 | -29.8493
20.022 | -35.2866 | -29.8557
20.0281 | -35.2871 | -29.8625
20.0342 | -35.2875 | -29.8693
20.0403 | -35.2879 | -29.8762
20.0464 | -35.2884 | -29.8836
20.0525 | -35.2889 | -29.89
20.0586 | -35.2893 | -29.8963
20.0647 | -35.2898 | -29.9027
20.0708 | -35.2902 | -29.909
20.0769 | -35.2906 | -29.9156
20.083 | -35.291 | -29.9222
20.0891 | -35.2914 | -29.9285
20.0952 | -35.2919 | -29.9351
20.1013 | -35.2923 | -29.942
20.1074 | -35.2928 | -29.9483
20.1135 | -35.2932 | -29.9546
20.1196 | -35.2936 | -29.9614
20.1257 | -35.294 | -29.9683
20.1319 | -35.2945 | -29.9749
20.138 | -35.2949 | -29.9813
20.1441 | -35.2954 | -29.9882
20.1502 | -35.2959 | -29.9952
20.1563 | -35.2963 | -30.0018
20.1624 | -35.2967 | -30.0085
20.1685 | -35.2971 | -30.0151
20.1746 | -35.2975 | -30.0214
20.1807 | -35.2979 | -30.0281
20.1868 | -35.2983 | -30.0348
20.1929 | -35.2986 | -30.0413
20.199 | -35.2991 | -30.0481
20.2051 | -35.2995 | -30.0545
20.2112 | -35.3 | -30.0614
20.2173 | -35.3004 | -30.0678
20.2234 | -35.3009 | -30.0746
20.2295 | -35.3013 | -30.0804
20.2356 | -35.3018 | -30.0869
20.2417 | -35.3023 | -30.0934
20.2478 | -35.3027 | -30.0997
20.2539 | -35.3032 | -30.1062
20.26 | -35.3036 | -30.1125
20.2661 | -35.304 | -30.1194
20.2723 | -35.3044 | -30.1261
20.2784 | -35.3048 | -30.1328
20.2845 | -35.3052 | -30.139
20.2906 | -35.3056 | -30.1457
20.2967 | -35.306 | -30.1522
20.3028 | -35.3064 | -30.1589
20.3089 | -35.3069 | -30.1655
20.315 | -35.3074 | -30.1724
20.3211 | -35.3079 | -30.1794
20.3272 | -35.3083 | -30.1863
20.3333 | -35.3086 | -30.193
20.3394 | -35.309 | -30.2001
20.3455 | -35.3094 | -30.2066
20.3516 | -35.3098 | -30.2134
20.3577 | -35.3102 | -30.2198
20.3638 | -35.3106 | -30.2265
20.3699 | -35.3109 | -30.2327
20.376 | -35.3114 | -30.2389
20.3821 | -35.3118 | -30.2458
20.3882 | -35.3122 | -30.2521
20.3943 | -35.3127 | -30.2585
20.4004 | -35.3132 | -30.2651
20.4065 | -35.3136 | -30.2716
20.4126 | -35.314 | -30.2784
20.4188 | -35.3144 | -30.2852
20.4249 | -35.3148 | -30.292
20.431 | -35.3153 | -30.2981
20.4371 | -35.3156 | -30.3052
20.4432 | -35.316 | -30.3121
20.4493 | -35.3164 | -30.3181
20.4554 | -35.3169 | -30.3248
20.4615 | -35.3173 | -30.3316
20.4676 | -35.3176 | -30.3386
20.4737 | -35.3181 | -30.3443
20.4798 | -35.3186 | -30.3511
20.4859 | -35.319 | -30.3575
20.492 | -35.3195 | -30.3645
20.4981 | -35.32 | -30.3716
20.5042 | -35.3204 | -30.3789
20.5103 | -35.3208 | -30.3848
20.5164 | -35.3212 | -30.391
20.5225 | -35.3216 | -30.3971
20.5286 | -35.322 | -30.4035
20.5347 | -35.3225 | -30.4104
20.5408 | -35.3228 | -30.4166
20.5469 | -35.3233 | -30.4232
20.553 | -35.3237 | -30.4302
20.5592 | -35.3241 | -30.4373
20.5653 | -35.3245 | -30.4441
20.5714 | -35.3248 | -30.4505
20.5775 | -35.3252 | -30.4575
20.5836 | -35.3257 | -30.4641
20.5897 | -35.3261 | -30.4708
20.5958 | -35.3265 | -30.4777
20.6019 | -35.3271 | -30.4847
20.608 | -35.3276 | -30.4907
20.6141 | -35.3279 | -30.4977
20.6202 | -35.3285 | -30.5043
20.6263 | -35.3289 | -30.511
20.6324 | -35.3294 | -30.5177
20.6385 | -35.3298 | -30.5241
20.6446 | -35.3302 | -30.5307
20.6507 | -35.3305 | -30.5367
20.6568 | -35.331 | -30.5435
20.6629 | -35.3314 | -30.5496
20.669 | -35.3318 | -30.5564
20.6751 | -35.3322 | -30.5632
20.6812 | -35.3326 | -30.5692
20.6873 | -35.3331 | -30.5759
20.6934 | -35.3335 | -30.582
20.6995 | -35.3339 | -30.589
20.7057 | -35.3344 | -30.5956
20.7118 | -35.335 | -30.6012
20.7179 | -35.3354 | -30.6074
20.724 | -35.3358 | -30.6137
20.7301 | -35.3362 | -30.6195
20.7362 | -35.3365 | -30.6264
20.7423 | -35.337 | -30.6325
20.7484 | -35.3373 | -30.638
20.7545 | -35.3378 | -30.6442
20.7606 | -35.3383 | -30.6509
20.7667 | -35.3387 | -30.657
20.7728 | -35.339 | -30.6636
20.7789 | -35.3395 | -30.6704
20.785 | -35.3398 | -30.6765
20.7911 | -35.3402 | -30.6828
20.7972 | -35.3406 | -30.6892
20.8033 | -35.3409 | -30.6952
20.8094 | -35.3414 | -30.7016
20.8155 | -35.3419 | -30.7081
20.8216 | -35.3422 | -30.714
20.8277 | -35.3428 | -30.7209
20.8338 | -35.3433 | -30.7272
20.8399 | -35.3437 | -30.7343
20.8461 | -35.344 | -30.7403
20.8522 | -35.3446 | -30.747
20.8583 | -35.3452 | -30.7538
20.8644 | -35.3458 | -30.761
20.8705 | -35.3464 | -30.7673
20.8766 | -35.3468 | -30.7738
20.8827 | -35.3472 | -30.7802
20.8888 | -35.3476 | -30.7874
20.8949 | -35.3481 | -30.7945
20.901 | -35.3486 | -30.8014
20.9071 | -35.349 | -30.8079
20.9132 | -35.3494 | -30.8146
20.9193 | -35.3499 | -30.8224
20.9254 | -35.3503 | -30.8299
20.9315 | -35.3509 | -30.8373
20.9376 | -35.3513 | -30.8439
20.9437 | -35.3516 | -30.8502
20.9498 | -35.3521 | -30.8573
20.9559 | -35.3525 | -30.8635
20.962 | -35.3529 | -30.8702
20.9681 | -35.3536 | -30.8765
20.9742 | -35.3541 | -30.8828
20.9803 | -35.3545 | -30.8894
20.9864 | -35.3549 | -30.8962
20.9926 | -35.3553 | -30.9017
20.9987 | -35.3557 | -30.9092
21.0048 | -35.3562 | -30.9159
21.0109 | -35.3565 | -30.9213
21.017 | -35.3571 | -30.9278
21.0231 | -35.3575 | -30.9349
21.0292 | -35.3579 | -30.9407
21.0353 | -35.3586 | -30.9467
21.0414 | -35.3588 | -30.9525
21.0475 | -35.3593 | -30.9597
21.0536 | -35.3597 | -30.9659
21.0597 | -35.36 | -30.9722
21.0658 | -35.3604 | -30.9788
21.0719 | -35.3608 | -30.9863
21.078 | -35.3612 | -30.992
21.0841 | -35.3615 | -30.9983
21.0902 | -35.362 | -31.005
21.0963 | -35.3624 | -31.0126
21.1024 | -35.3629 | -31.0183
21.1085 | -35.3637 | -31.0247
21.1146 | -35.3641 | -31.0305
21.1207 | -35.3646 | -31.037
21.1268 | -35.3651 | -31.0433
21.133 | -35.3657 | -31.0492
21.1391 | -35.3662 | -31.0553
21.1452 | -35.3666 | -31.0623
21.1513 | -35.367 | -31.0672
21.1574 | -35.3673 | -31.0747
21.1635 | -35.3678 | -31.0817
21.1696 | -35.3682 | -31.0883
21.1757 | -35.3686 | -31.0951
21.1818 | -35.3691 | -31.1022
21.1879 | -35.3696 | -31.1079
21.194 | -35.37 | -31.1145
21.2001 | -35.3704 | -31.121
21.2062 | -35.3707 | -31.1275
21.2123 | -35.3713 | -31.1335
21.2184 | -35.3718 | -31.1398
21.2245 | -35.3722 | -31.1463
21.2306 | -35.3726 | -31.1518
21.2367 | -35.3731 | -31.1584
21.2428 | -35.3735 | -31.1639
21.2489 | -35.3739 | -31.1702
21.255 | -35.3743 | -31.1763
21.2611 | -35.3748 | -31.1827
21.2672 | -35.3751 | -31.1891
21.2733 | -35.3756 | -31.1954
21.2795 | -35.376 | -31.2026
21.2856 | -35.3766 | -31.2082
21.2917 | -35.3768 | -31.2142
21.2978 | -35.3774 | -31.2206
21.3039 | -35.3779 | -31.2286
21.31 | -35.3784 | -31.2358
21.3161 | -35.3787 | -31.2417
21.3222 | -35.3791 | -31.2489
21.3283 | -35.3795 | -31.2558
21.3344 | -35.38 | -31.2616
21.3405 | -35.3807 | -31.2675
21.3466 | -35.3811 | -31.2746
21.3527 | -35.3816 | -31.2806
21.3588 | -35.3821 | -31.2872
21.3649 | -35.3826 | -31.2945
21.371 | -35.3829 | -31.3014
21.3771 | -35.3834 | -31.3082
21.3832 | -35.3839 | -31.3144
21.3893 | -35.3844 | -31.32
21.3954 | -35.3848 | -31.3271
21.4015 | -35.3853 | -31.3334
21.4076 | -35.3857 | -31.3393
21.4137 | -35.3862 | -31.3453
21.4199 | -35.3866 | -31.3515
21.426 | -35.3872 | -31.359
21.4321 | -35.3877 | -31.3653
21.4382 | -35.3881 | -31.3723
21.4443 | -35.3886 | -31.3789
21.4504 | -35.3893 | -31.3859
21.4565 | -35.3895 | -31.3935
21.4626 | -35.39 | -31.4
21.4687 | -35.3903 | -31.4061
21.4748 | -35.3908 | -31.4126
21.4809 | -35.3914 | -31.4193
21.487 | -35.3919 | -31.4264
21.4931 | -35.3923 | -31.4331
21.4992 | -35.3928 | -31.4399
21.5053 | -35.3933 | -31.4462
21.5114 | -35.3937 | -31.4526
21.5175 | -35.3941 | -31.4598
21.5236 | -35.3943 | -31.4667
21.5297 | -35.3948 | -31.4735
21.5358 | -35.3954 | -31.4796
21.5419 | -35.3959 | -31.4867
21.548 | -35.3964 | -31.4933
21.5541 | -35.3967 | -31.5002
21.5602 | -35.3972 | -31.5059
21.5664 | -35.3977 | -31.513
21.5725 | -35.3982 | -31.5186
21.5786 | -35.3987 | -31.525
21.5847 | -35.399 | -31.5317
21.5908 | -35.3995 | -31.5382
21.5969 | -35.4 | -31.5466
21.603 | -35.4002 | -31.5528
21.6091 | -35.4008 | -31.5601
21.6152 | -35.4014 | -31.566
21.6213 | -35.4018 | -31.5729
21.6274 | -35.4023 | -31.5792
21.6335 | -35.4027 | -31.5838
21.6396 | -35.4033 | -31.5901
21.6457 | -35.4038 | -31.5965
21.6518 | -35.4042 | -31.6029
21.6579 | -35.4047 | -31.6096
21.664 | -35.4052 | -31.6166
21.6701 | -35.4056 | -31.6225
21.6762 | -35.406 | -31.6295
21.6823 | -35.4063 | -31.6367
21.6884 | -35.4066 | -31.6426
21.6945 | -35.407 | -31.6499
21.7006 | -35.4074 | -31.6566
21.7068 | -35.4076 | -31.6629
21.7129 | -35.4081 | -31.6688
21.719 | -35.4085 | -31.6768
21.7251 | -35.409 | -31.6827
21.7312 | -35.4095 | -31.69
21.7373 | -35.41 | -31.6977
21.7434 | -35.4102 | -31.704
21.7495 | -35.4108 | -31.7109
21.7556 | -35.4112 | -31.7175
21.7617 | -35.4117 | -31.7257
21.7678 | -35.4124 | -31.7321
21.7739 | -35.4128 | -31.7395
21.78 | -35.4132 | -31.7464
21.7861 | -35.4135 | -31.7519
21.7922 | -35.4142 | -31.7587
21.7983 | -35.4146 | -31.7648
21.8044 | -35.4152 | -31.7713
21.8105 | -35.4156 | -31.7787
21.8166 | -35.4161 | -31.7864
21.8227 | -35.4167 | -31.7934
21.8288 | -35.4171 | -31.7984
21.8349 | -35.4177 | -31.8048
21.841 | -35.4181 | -31.8099
21.8471 | -35.4188 | -31.8182
21.8533 | -35.4191 | -31.8254
21.8594 | -35.4194 | -31.8329
21.8655 | -35.4199 | -31.8393
21.8716 | -35.4203 | -31.8452
21.8777 | -35.4205 | -31.8519
21.8838 | -35.421 | -31.8581
21.8899 | -35.4215 | -31.8647
21.896 | -35.4218 | -31.8714
21.9021 | -35.4224 | -31.8782
21.9082 | -35.4229 | -31.885
21.9143 | -35.4235 | -31.8917
21.9204 | -35.4241 | -31.898
21.9265 | -35.4246 | -31.9052
21.9326 | -35.4251 | -31.9116
21.9387 | -35.4255 | -31.918
21.9448 | -35.426 | -31.9247
21.9509 | -35.4263 | -31.9313
21.957 | -35.4268 | -31.9381
21.9631 | -35.4273 | -31.9444
21.9692 | -35.4278 | -31.9511
21.9753 | -35.4282 | -31.9577
21.9814 | -35.4287 | -31.964
21.9875 | -35.4293 | -31.9701
21.9937 | -35.4298 | -31.9765
21.9998 | -35.4303 | -31.9836
22.0059 | -35.4308 | -31.9898
22.012 | -35.4312 | -31.9966
22.0181 | -35.4317 | -32.0037
22.0242 | -35.432 | -32.0103
22.0303 | -35.4324 | -32.0175
22.0364 | -35.4329 | -32.0236
22.0425 | -35.4334 | -32.0306
22.0486 | -35.4339 | -32.0374
22.0547 | -35.4343 | -32.0439
22.0608 | -35.4346 | -32.0495
22.0669 | -35.4353 | -32.056
22.073 | -35.4357 | -32.0625
22.0791 | -35.4363 | -32.069
22.0852 | -35.4366 | -32.0757
22.0913 | -35.4372 | -32.0825
22.0974 | -35.4376 | -32.0885
22.1035 | -35.4381 | -32.0947
22.1096 | -35.4386 | -32.101
22.1157 | -35.4391 | -32.1077
22.1218 | -35.4396 | -32.1144
22.1279 | -35.4403 | -32.1201
22.134 | -35.4405 | -32.1265
22.1402 | -35.4409 | -32.1337
22.1463 | -35.4413 | -32.1407
22.1524 | -35.4418 | -32.1467
22.1585 | -35.4422 | -32.1536
22.1646 | -35.4427 | -32.1605
22.1707 | -35.4432 | -32.1678
22.1768 | -35.4438 | -32.1746
22.1829 | -35.4442 | -32.1803
22.189 | -35.4447 | -32.1865
22.1951 | -35.4451 | -32.1935
22.2012 | -35.4456 | -32.2005
22.2073 | -35.4461 | -32.2069
22.2134 | -35.4464 | -32.2135
22.2195 | -35.4468 | -32.2196
22.2256 | -35.4473 | -32.2257
22.2317 | -35.4478 | -32.2332
22.2378 | -35.4482 | -32.2402
22.2439 | -35.4487 | -32.2471
22.25 | -35.4491 | -32.2534
22.2561 | -35.4495 | -32.2602
22.2622 | -35.4501 | -32.2663
22.2683 | -35.4504 | -32.2744
22.2744 | -35.451 | -32.2805
22.2806 | -35.4514 | -32.2872
22.2867 | -35.4519 | -32.2945
22.2928 | -35.4524 | -32.3011
22.2989 | -35.4529 | -32.3085
22.305 | -35.4533 | -32.3152
22.3111 | -35.4537 | -32.322
22.3172 | -35.4542 | -32.3288
22.3233 | -35.4548 | -32.3352
22.3294 | -35.4552 | -32.3418
22.3355 | -35.4557 | -32.3488
22.3416 | -35.4561 | -32.3549
22.3477 | -35.4566 | -32.3615
22.3538 | -35.457 | -32.3679
22.3599 | -35.4574 | -32.3743
22.366 | -35.458 | -32.3807
22.3721 | -35.4586 | -32.3874
22.3782 | -35.459 | -32.3935
22.3843 | -35.4596 | -32.3996
22.3904 | -35.46 | -32.4062
22.3965 | -35.4605 | -32.4132
22.4026 | -35.461 | -32.4203
22.4087 | -35.4616 | -32.4281
22.4148 | -35.462 | -32.4353
22.4209 | -35.4624 | -32.4417
22.4271 | -35.4629 | -32.4486
22.4332 | -35.4634 | -32.455
22.4393 | -35.4639 | -32.4617
22.4454 | -35.4646 | -32.4675
22.4515 | -35.4651 | -32.4732
22.4576 | -35.4655 | -32.4797
22.4637 | -35.466 | -32.487
22.4698 | -35.4665 | -32.4934
22.4759 | -35.4671 | -32.4991
22.482 | -35.4676 | -32.5054
22.4881 | -35.4681 | -32.5127
22.4942 | -35.4686 | -32.5196
22.5003 | -35.4689 | -32.5266
22.5064 | -35.4694 | -32.5327
22.5125 | -35.47 | -32.5386
22.5186 | -35.4705 | -32.5456
22.5247 | -35.471 | -32.552
22.5308 | -35.4716 | -32.5591
22.5369 | -35.472 | -32.566
22.543 | -35.4725 | -32.5724
22.5491 | -35.4729 | -32.5791
22.5552 | -35.4734 | -32.5855
22.5613 | -35.4739 | -32.5921
22.5675 | -35.4744 | -32.5989
22.5736 | -35.4749 | -32.6056
22.5797 | -35.4755 | -32.6125
22.5858 | -35.4761 | -32.6185
22.5919 | -35.4766 | -32.6251
22.598 | -35.477 | -32.6317
22.6041 | -35.4776 | -32.6388
22.6102 | -35.4781 | -32.6454
22.6163 | -35.4787 | -32.6521
22.6224 | -35.4792 | -32.6585
22.6285 | -35.4797 | -32.6649
22.6346 | -35.4801 | -32.6718
22.6407 | -35.4807 | -32.6781
22.6468 | -35.4811 | -32.6854
22.6529 | -35.4817 | -32.6918
22.659 | -35.4821 | -32.6984
22.6651 | -35.4826 | -32.7055
22.6712 | -35.4831 | -32.7124
22.6773 | -35.4836 | -32.7188
22.6834 | -35.4842 | -32.7259
22.6895 | -35.4846 | -32.7325
22.6956 | -35.4851 | -32.7394
22.7017 | -35.4856 | -32.7459
22.7079 | -35.486 | -32.7524
22.714 | -35.4864 | -32.7592
22.7201 | -35.4869 | -32.7664
22.7262 | -35.4873 | -32.7737
22.7323 | -35.4877 | -32.7806
22.7384 | -35.4881 | -32.7874
22.7445 | -35.4886 | -32.7943
22.7506 | -35.4891 | -32.8006
22.7567 | -35.4895 | -32.8075
22.7628 | -35.4899 | -32.8138
22.7689 | -35.4904 | -32.8202
22.775 | -35.4909 | -32.8267
22.7811 | -35.4914 | -32.8333
22.7872 | -35.4919 | -32.8405
22.7933 | -35.4924 | -32.846
22.7994 | -35.4929 | -32.8524
22.8055 | -35.4934 | -32.8585
22.8116 | -35.4939 | -32.865
22.8177 | -35.4944 | -32.8717
22.8238 | -35.4949 | -32.878
22.8299 | -35.4954 | -32.8848
22.836 | -35.4959 | -32.8917
22.8421 | -35.4964 | -32.8978
22.8482 | -35.4968 | -32.9045
22.8544 | -35.4974 | -32.9113
22.8605 | -35.4979 | -32.9175
22.8666 | -35.4983 | -32.9244
22.8727 | -35.4988 | -32.9303
22.8788 | -35.4993 | -32.937
22.8849 | -35.4998 | -32.9437
22.891 | -35.5002 | -32.9502
22.8971 | -35.5007 | -32.957
22.9032 | -35.5012 | -32.9627
22.9093 | -35.5017 | -32.9694
22.9154 | -35.5022 | -32.9759
22.9215 | -35.5026 | -32.983
22.9276 | -35.503 | -32.9898
22.9337 | -35.5035 | -32.9973
22.9398 | -35.504 | -33.0036
22.9459 | -35.5046 | -33.01
22.952 | -35.505 | -33.017
22.9581 | -35.5055 | -33.0237
22.9642 | -35.5058 | -33.0309
22.9703 | -35.5064 | -33.0372
22.9764 | -35.5069 | -33.0439
22.9825 | -35.5074 | -33.0497
22.9886 | -35.5078 | -33.0563
22.9948 | -35.5083 | -33.0631
23.0009 | -35.5089 | -33.0692
23.007 | -35.5094 | -33.0763
23.0131 | -35.5099 | -33.0836
23.0192 | -35.5105 | -33.09
23.0253 | -35.511 | -33.0967
23.0314 | -35.5115 | -33.1037
23.0375 | -35.5119 | -33.1108
23.0436 | -35.5124 | -33.1176
23.0497 | -35.513 | -33.1244
23.0558 | -35.5134 | -33.1312
23.0619 | -35.5139 | -33.1382
23.068 | -35.5144 | -33.1451
23.0741 | -35.515 | -33.1514
23.0802 | -35.5155 | -33.1578
23.0863 | -35.5159 | -33.1646
23.0924 | -35.5164 | -33.1709
23.0985 | -35.517 | -33.1774
23.1046 | -35.5174 | -33.1841
23.1107 | -35.5179 | -33.1905
23.1168 | -35.5183 | -33.1974
23.1229 | -35.5187 | -33.2038
23.129 | -35.5193 | -33.2109
23.1351 | -35.5198 | -33.2175
23.1413 | -35.5202 | -33.2242
23.1474 | -35.5207 | -33.2306
23.1535 | -35.5213 | -33.2371
23.1596 | -35.5218 | -33.2433
23.1657 | -35.5223 | -33.2502
23.1718 | -35.5227 | -33.2569
23.1779 | -35.5233 | -33.2637
23.184 | -35.5238 | -33.2707
23.1901 | -35.5243 | -33.2771
23.1962 | -35.5249 | -33.2834
23.2023 | -35.5253 | -33.2905
23.2084 | -35.5258 | -33.2972
23.2145 | -35.5263 | -33.3037
23.2206 | -35.5268 | -33.3101
23.2267 | -35.5272 | -33.317
23.2328 | -35.5277 | -33.3241
23.2389 | -35.5283 | -33.3302
23.245 | -35.5288 | -33.3367
23.2511 | -35.5293 | -33.3427
23.2572 | -35.5299 | -33.3487
23.2633 | -35.5304 | -33.3555
23.2694 | -35.5308 | -33.3622
23.2755 | -35.5312 | -33.369
23.2817 | -35.5316 | -33.3755
23.2878 | -35.5322 | -33.3826
23.2939 | -35.5328 | -33.3889
23.3 | -35.5332 | -33.3958
23.3061 | -35.5337 | -33.4019
23.3122 | -35.5343 | -33.4086
23.3183 | -35.5348 | -33.4151
23.3244 | -35.5353 | -33.4213
23.3305 | -35.5359 | -33.4281
23.3366 | -35.5364 | -33.4342
23.3427 | -35.5369 | -33.4405
23.3488 | -35.5375 | -33.4478
23.3549 | -35.5379 | -33.4541
23.361 | -35.5384 | -33.4606
23.3671 | -35.5389 | -33.4666
23.3732 | -35.5393 | -33.4729
23.3793 | -35.5398 | -33.4797
23.3854 | -35.5404 | -33.4865
23.3915 | -35.5409 | -33.492
23.3976 | -35.5413 | -33.499
23.4037 | -35.5418 | -33.5052
23.4098 | -35.5424 | -33.5113
23.4159 | -35.5429 | -33.5178
23.422 | -35.5434 | -33.5246
23.4282 | -35.5439 | -33.5307
23.4343 | -35.5444 | -33.5375
23.4404 | -35.5449 | -33.5443
23.4465 | -35.5455 | -33.5503
23.4526 | -35.546 | -33.5567
23.4587 | -35.5465 | -33.5637
23.4648 | -35.547 | -33.5699
23.4709 | -35.5474 | -33.5763
23.477 | -35.548 | -33.583
23.4831 | -35.5485 | -33.5896
23.4892 | -35.549 | -33.5957
23.4953 | -35.5494 | -33.6023
23.5014 | -35.55 | -33.6086
23.5075 | -35.5505 | -33.6147
23.5136 | -35.551 | -33.6214
23.5197 | -35.5514 | -33.6271
23.5258 | -35.5521 | -33.6333
23.5319 | -35.5526 | -33.6392
23.538 | -35.5531 | -33.6458
23.5441 | -35.5536 | -33.652
23.5502 | -35.554 | -33.6583
23.5563 | -35.5545 | -33.665
23.5624 | -35.555 | -33.6712
23.5686 | -35.5556 | -33.6777
23.5747 | -35.5561 | -33.6846
23.5808 | -35.5565 | -33.6911
23.5869 | -35.557 | -33.6973
23.593 | -35.5576 | -33.7036
23.5991 | -35.5581 | -33.7105
23.6052 | -35.5587 | -33.7169
23.6113 | -35.5592 | -33.723
23.6174 | -35.5596 | -33.7294
23.6235 | -35.5601 | -33.7358
23.6296 | -35.5606 | -33.7419
23.6357 | -35.5611 | -33.7487
23.6418 | -35.5616 | -33.7556
23.6479 | -35.562 | -33.7621
23.654 | -35.5626 | -33.7685
23.6601 | -35.5631 | -33.7754
23.6662 | -35.5636 | -33.7814
23.6723 | -35.5643 | -33.7874
23.6784 | -35.5648 | -33.7942
23.6845 | -35.5653 | -33.8016
23.6906 | -35.5656 | -33.808
23.6967 | -35.5661 | -33.8146
23.7028 | -35.5667 | -33.8213
23.7089 | -35.5671 | -33.8275
23.7151 | -35.5677 | -33.8341
23.7212 | -35.5681 | -33.8406
23.7273 | -35.5686 | -33.8472
23.7334 | -35.569 | -33.8538
23.7395 | -35.5695 | -33.8607
23.7456 | -35.57 | -33.8661
23.7517 | -35.5706 | -33.8728
23.7578 | -35.5711 | -33.8789
23.7639 | -35.5715 | -33.8855
23.77 | -35.572 | -33.8916
23.7761 | -35.5724 | -33.8982
23.7822 | -35.573 | -33.904
23.7883 | -35.5734 | -33.9109
23.7944 | -35.5739 | -33.9168
23.8005 | -35.5744 | -33.9231
23.8066 | -35.575 | -33.9297
23.8127 | -35.5755 | -33.936
23.8188 | -35.5759 | -33.9426
23.8249 | -35.5764 | -33.9492
23.831 | -35.5769 | -33.9558
23.8371 | -35.5773 | -33.9623
23.8432 | -35.5779 | -33.9685
23.8493 | -35.5784 | -33.9748
23.8555 | -35.5788 | -33.9809
23.8616 | -35.5793 | -33.9873
23.8677 | -35.5798 | -33.9936
23.8738 | -35.5803 | -33.9998
23.8799 | -35.5808 | -34.0056
23.886 | -35.5813 | -34.0113
23.8921 | -35.5818 | -34.0177
23.8982 | -35.5823 | -34.0242
23.9043 | -35.5828 | -34.0312
23.9104 | -35.5833 | -34.0369
23.9165 | -35.5838 | -34.0432
23.9226 | -35.5843 | -34.0494
23.9287 | -35.5849 | -34.0562
23.9348 | -35.5852 | -34.0622
23.9409 | -35.5858 | -34.0689
23.947 | -35.5862 | -34.0756
23.9531 | -35.5868 | -34.0804
23.9592 | -35.5873 | -34.0858
23.9653 | -35.5877 | -34.0926
23.9714 | -35.5883 | -34.099
23.9775 | -35.5888 | -34.1053
23.9836 | -35.5893 | -34.1115
23.9897 | -35.5897 | -34.1176
23.9958 | -35.5901 | -34.1236
24.002 | -35.5906 | -34.1302
24.0081 | -35.5911 | -34.1365
24.0142 | -35.5916 | -34.1426
24.0203 | -35.5921 | -34.1485
24.0264 | -35.5925 | -34.1551
24.0325 | -35.5931 | -34.1612
24.0386 | -35.5937 | -34.1678
24.0447 | -35.5942 | -34.1739
24.0508 | -35.5947 | -34.1801
24.0569 | -35.5951 | -34.1859
24.063 | -35.5956 | -34.1917
24.0691 | -35.596 | -34.1975
24.0752 | -35.5965 | -34.2037
24.0813 | -35.5971 | -34.2099
24.0874 | -35.5975 | -34.2167
24.0935 | -35.598 | -34.2237
24.0996 | -35.5985 | -34.2294
24.1057 | -35.599 | -34.2361
24.1118 | -35.5994 | -34.2425
24.1179 | -35.6 | -34.249
24.124 | -35.6005 | -34.2553
24.1301 | -35.6009 | -34.2613
24.1362 | -35.6013 | -34.2678
24.1424 | -35.6018 | -34.2738
24.1485 | -35.6023 | -34.2806
24.1546 | -35.6028 | -34.287
24.1607 | -35.6033 | -34.2931
24.1668 | -35.6038 | -34.2984
24.1729 | -35.6042 | -34.3046
24.179 | -35.6047 | -34.3108
24.1851 | -35.6052 | -34.3175
24.1912 | -35.6056 | -34.3236
24.1973 | -35.606 | -34.3297
24.2034 | -35.6064 | -34.3349
24.2095 | -35.6069 | -34.3411
24.2156 | -35.6074 | -34.3475
24.2217 | -35.6078 | -34.3536
24.2278 | -35.6083 | -34.3601
24.2339 | -35.6087 | -34.3656
24.24 | -35.6091 | -34.371
24.2461 | -35.6096 | -34.3771
24.2522 | -35.6101 | -34.3842
24.2583 | -35.6105 | -34.3902
24.2644 | -35.6109 | -34.3964
24.2705 | -35.6114 | -34.4024
24.2766 | -35.6119 | -34.4086
24.2827 | -35.6124 | -34.4143
24.2889 | -35.6129 | -34.4202
24.295 | -35.6134 | -34.4264
24.3011 | -35.6138 | -34.4325
24.3072 | -35.6143 | -34.438
24.3133 | -35.6147 | -34.4443
24.3194 | -35.6151 | -34.45
24.3255 | -35.6155 | -34.4554
24.3316 | -35.616 | -34.462
24.3377 | -35.6163 | -34.4684
24.3438 | -35.6168 | -34.4746
24.3499 | -35.6173 | -34.4805
24.356 | -35.6178 | -34.4865
24.3621 | -35.6184 | -34.4922
24.3682 | -35.6189 | -34.4981
24.3743 | -35.6193 | -34.5043
24.3804 | -35.6197 | -34.5093
24.3865 | -35.6201 | -34.5156
24.3926 | -35.6205 | -34.5209
24.3987 | -35.6211 | -34.5275
24.4048 | -35.6216 | -34.5335
24.4109 | -35.6221 | -34.5397
24.417 | -35.6226 | -34.5456
24.4231 | -35.6231 | -34.5522
24.4293 | -35.6236 | -34.5582
24.4354 | -35.624 | -34.5642
24.4415 | -35.6245 | -34.5697
24.4476 | -35.625 | -34.5754
24.4537 | -35.6255 | -34.5816
24.4598 | -35.6259 | -34.5874
24.4659 | -35.6264 | -34.5939
24.472 | -35.6267 | -34.599
24.4781 | -35.6272 | -34.6053
24.4842 | -35.6277 | -34.6112
24.4903 | -35.6281 | -34.6179
24.4964 | -35.6286 | -34.6235
24.5025 | -35.629 | -34.6296
24.5086 | -35.6294 | -34.6355
24.5147 | -35.6299 | -34.6415
24.5208 | -35.6303 | -34.6467
24.5269 | -35.6308 | -34.6531
24.533 | -35.6313 | -34.6587
24.5391 | -35.6317 | -34.6645
24.5452 | -35.6322 | -34.6707
24.5513 | -35.6326 | -34.6762
24.5574 | -35.6332 | -34.682
24.5635 | -35.6336 | -34.6882
24.5696 | -35.6341 | -34.6942
24.5758 | -35.6347 | -34.7004
24.5819 | -35.6352 | -34.7059
24.588 | -35.6356 | -34.7111
24.5941 | -35.6361 | -34.7171
24.6002 | -35.6365 | -34.7236
24.6063 | -35.6369 | -34.7299
24.6124 | -35.6374 | -34.7357
24.6185 | -35.6379 | -34.7418
24.6246 | -35.6383 | -34.7469
24.6307 | -35.6389 | -34.7524
24.6368 | -35.6394 | -34.7587
24.6429 | -35.6398 | -34.7647
24.649 | -35.6403 | -34.7706
24.6551 | -35.6408 | -34.7764
24.6612 | -35.6413 | -34.7813
24.6673 | -35.642 | -34.7877
24.6734 | -35.6425 | -34.7939
24.6795 | -35.6429 | -34.7994
24.6856 | -35.6434 | -34.8053
24.6917 | -35.6439 | -34.8118
24.6978 | -35.6444 | -34.8172
24.7039 | -35.6449 | -34.8235
24.71 | -35.6453 | -34.8291
24.7162 | -35.6458 | -34.8343
24.7223 | -35.6462 | -34.8402
24.7284 | -35.6466 | -34.8454
24.7345 | -35.6471 | -34.851
24.7406 | -35.6475 | -34.8574
24.7467 | -35.6479 | -34.8633
24.7528 | -35.6483 | -34.869
24.7589 | -35.6488 | -34.8754
24.765 | -35.6492 | -34.8796
24.7711 | -35.6497 | -34.8846
24.7772 | -35.6501 | -34.89
24.7833 | -35.6505 | -34.8959
24.7894 | -35.651 | -34.9018
24.7955 | -35.6515 | -34.9072
24.8016 | -35.652 | -34.9131
24.8077 | -35.6524 | -34.9193
24.8138 | -35.6529 | -34.9252
24.8199 | -35.6533 | -34.9315
24.826 | -35.6537 | -34.9373
24.8321 | -35.6541 | -34.9433
24.8382 | -35.6545 | -34.9483
24.8443 | -35.6549 | -34.954
24.8504 | -35.6554 | -34.9604
24.8565 | -35.6559 | -34.9665
24.8627 | -35.6563 | -34.9721
24.8688 | -35.6567 | -34.9775
24.8749 | -35.6573 | -34.983
24.881 | -35.6577 | -34.9883
24.8871 | -35.6581 | -34.9938
24.8932 | -35.6586 | -34.9995
24.8993 | -35.659 | -35.0054
24.9054 | -35.6594 | -35.0101
24.9115 | -35.6599 | -35.0156
24.9176 | -35.6604 | -35.0214
24.9237 | -35.6608 | -35.0272
24.9298 | -35.6612 | -35.0329
24.9359 | -35.6617 | -35.0377
24.942 | -35.6621 | -35.0426
24.9481 | -35.6624 | -35.0483
24.9542 | -35.6629 | -35.0537
24.9603 | -35.6633 | -35.0594
24.9664 | -35.6637 | -35.0648
24.9725 | -35.6642 | -35.07
24.9786 | -35.6646 | -35.0748
24.9847 | -35.6651 | -35.0806
24.9908 | -35.6655 | -35.086
24.9969 | -35.6659 | -35.092
25.0031 | -35.6664 | -35.0977
25.0092 | -35.6668 | -35.1032
25.0153 | -35.6672 | -35.1088
25.0214 | -35.6677 | -35.1147
25.0275 | -35.6681 | -35.12
25.0336 | -35.6685 | -35.1259
25.0397 | -35.6689 | -35.1311
25.0458 | -35.6693 | -35.136
25.0519 | -35.6697 | -35.1414
25.058 | -35.6701 | -35.1464
25.0641 | -35.6705 | -35.1517
25.0702 | -35.6709 | -35.1578
25.0763 | -35.6713 | -35.1633
25.0824 | -35.6718 | -35.1686
25.0885 | -35.6721 | -35.1741
25.0946 | -35.6726 | -35.1798
25.1007 | -35.673 | -35.1857
25.1068 | -35.6735 | -35.1915
25.1129 | -35.6739 | -35.1975
25.119 | -35.6744 | -35.2019
25.1251 | -35.6749 | -35.2077
25.1312 | -35.6753 | -35.2134
25.1373 | -35.6757 | -35.2195
25.1435 | -35.6761 | -35.2254
25.1496 | -35.6765 | -35.2295
25.1557 | -35.6769 | -35.2356
25.1618 | -35.6773 | -35.2411
25.1679 | -35.6778 | -35.2466
25.174 | -35.6782 | -35.2523
25.1801 | -35.6787 | -35.2581
25.1862 | -35.6791 | -35.2631
25.1923 | -35.6796 | -35.2684
25.1984 | -35.68 | -35.2742
25.2045 | -35.6804 | -35.2793
25.2106 | -35.6808 | -35.2844
25.2167 | -35.6813 | -35.2901
25.2228 | -35.6815 | -35.2953
25.2289 | -35.682 | -35.3009
25.235 | -35.6824 | -35.3063
25.2411 | -35.6829 | -35.3116
25.2472 | -35.6832 | -35.3173
25.2533 | -35.6837 | -35.3226
25.2594 | -35.684 | -35.329
25.2655 | -35.6844 | -35.3348
25.2716 | -35.6848 | -35.34
25.2777 | -35.6852 | -35.3451
25.2838 | -35.6857 | -35.3506
25.29 | -35.6861 | -35.3562
25.2961 | -35.6864 | -35.3613
25.3022 | -35.6869 | -35.3666
25.3083 | -35.6873 | -35.3729
25.3144 | -35.6878 | -35.3779
25.3205 | -35.6882 | -35.3838
25.3266 | -35.6886 | -35.3892
25.3327 | -35.6889 | -35.3948
25.3388 | -35.6894 | -35.4009
25.3449 | -35.6898 | -35.4066
25.351 | -35.6903 | -35.4123
25.3571 | -35.6907 | -35.4171
25.3632 | -35.691 | -35.4224
25.3693 | -35.6913 | -35.4276
25.3754 | -35.6918 | -35.433
25.3815 | -35.6922 | -35.4387
25.3876 | -35.6926 | -35.444
25.3937 | -35.693 | -35.4493
25.3998 | -35.6934 | -35.4539
25.4059 | -35.6938 | -35.4595
25.412 | -35.6943 | -35.4649
25.4181 | -35.6949 | -35.4706
25.4242 | -35.6953 | -35.4762
25.4304 | -35.6957 | -35.4815
25.4365 | -35.696 | -35.4868
25.4426 | -35.6965 | -35.4923
25.4487 | -35.6969 | -35.4976
25.4548 | -35.6973 | -35.5028
25.4609 | -35.6978 | -35.508
25.467 | -35.6982 | -35.5124
25.4731 | -35.6986 | -35.5173
25.4792 | -35.6991 | -35.5226
25.4853 | -35.6995 | -35.5285
25.4914 | -35.7 | -35.5334
25.4975 | -35.7005 | -35.5385
25.5036 | -35.7009 | -35.5443
25.5097 | -35.7013 | -35.5499
25.5158 | -35.7017 | -35.5555
25.5219 | -35.702 | -35.5611
25.528 | -35.7025 | -35.5666
25.5341 | -35.7028 | -35.5715
25.5402 | -35.7032 | -35.5772
25.5463 | -35.7036 | -35.5827
25.5524 | -35.7041 | -35.588
25.5585 | -35.7045 | -35.5932
25.5646 | -35.705 | -35.5987
25.5707 | -35.7054 | -35.6041
25.5769 | -35.7057 | -35.609
25.583 | -35.7062 | -35.6138
25.5891 | -35.7067 | -35.6189
25.5952 | -35.7071 | -35.6245
25.6013 | -35.7075 | -35.6298
25.6074 | -35.7078 | -35.6339
25.6135 | -35.7083 | -35.6395
25.6196 | -35.7087 | -35.6446
25.6257 | -35.7091 | -35.65
25.6318 | -35.7094 | -35.6555
25.6379 | -35.7097 | -35.6605
25.644 | -35.7101 | -35.6655
25.6501 | -35.7105 | -35.6707
25.6562 | -35.7109 | -35.6765
25.6623 | -35.7113 | -35.6816
25.6684 | -35.7117 | -35.687
25.6745 | -35.712 | -35.6922
25.6806 | -35.7124 | -35.6964
25.6867 | -35.7129 | -35.7015
25.6928 | -35.7132 | -35.7068
25.6989 | -35.7136 | -35.712
25.705 | -35.7139 | -35.7173
25.7111 | -35.7144 | -35.7224
25.7173 | -35.7149 | -35.7276
25.7234 | -35.7153 | -35.7323
25.7295 | -35.7157 | -35.7375
25.7356 | -35.716 | -35.7425
25.7417 | -35.7164 | -35.7475
25.7478 | -35.7168 | -35.7512
25.7539 | -35.7172 | -35.7562
25.76 | -35.7176 | -35.7618
25.7661 | -35.718 | -35.7671
25.7722 | -35.7183 | -35.772
25.7783 | -35.7187 | -35.7768
25.7844 | -35.719 | -35.7817
25.7905 | -35.7193 | -35.7872
25.7966 | -35.7197 | -35.792
25.8027 | -35.72 | -35.7973
25.8088 | -35.7204 | -35.8024
25.8149 | -35.7208 | -35.8072
25.821 | -35.7212 | -35.812
25.8271 | -35.7215 | -35.8171
25.8332 | -35.722 | -35.8223
25.8393 | -35.7224 | -35.8276
25.8454 | -35.7228 | -35.8326
25.8515 | -35.7232 | -35.8375
25.8576 | -35.7236 | -35.8425
25.8638 | -35.724 | -35.848
25.8699 | -35.7244 | -35.8528
25.876 | -35.7247 | -35.8573
25.8821 | -35.7252 | -35.8622
25.8882 | -35.7255 | -35.8662
25.8943 | -35.7259 | -35.8718
25.9004 | -35.7262 | -35.8767
25.9065 | -35.7266 | -35.8818
25.9126 | -35.727 | -35.8868
25.9187 | -35.7273 | -35.8919
25.9248 | -35.7278 | -35.8968
25.9309 | -35.7281 | -35.9018
25.937 | -35.7285 | -35.9069
25.9431 | -35.7289 | -35.9122
25.9492 | -35.7292 | -35.9171
25.9553 | -35.7296 | -35.9224
25.9614 | -35.7299 | -35.9265
25.9675 | -35.7303 | -35.9318
25.9736 | -35.7308 | -35.9371
25.9797 | -35.7312 | -35.9428
25.9858 | -35.7316 | -35.948
25.9919 | -35.7319 | -35.9533
25.998 | -35.7323 | -35.9576
26.0042 | -35.7326 | -35.9627
26.0103 | -35.733 | -35.9675
26.0164 | -35.7333 | -35.9732
26.0225 | -35.7337 | -35.9783
26.0286 | -35.7342 | -35.9834
26.0347 | -35.7346 | -35.9882
26.0408 | -35.7351 | -35.9928
26.0469 | -35.7355 | -35.9973
26.053 | -35.7359 | -36.0025
26.0591 | -35.7363 | -36.0075
26.0652 | -35.7367 | -36.0114
26.0713 | -35.737 | -36.0167
26.0774 | -35.7374 | -36.0218
26.0835 | -35.7378 | -36.0268
26.0896 | -35.7381 | -36.0318
26.0957 | -35.7385 | -36.0365
26.1018 | -35.7388 | -36.0414
26.1079 | -35.7393 | -36.0462
26.114 | -35.7396 | -36.0512
26.1201 | -35.74 | -36.0559
26.1262 | -35.7404 | -36.0609
26.1323 | -35.7408 | -36.066
26.1384 | -35.7412 | -36.0703
26.1445 | -35.7414 | -36.0756
26.1507 | -35.7417 | -36.0807
26.1568 | -35.742 | -36.0856
26.1629 | -35.7424 | -36.0907
26.169 | -35.7428 | -36.0952
26.1751 | -35.7431 | -36.0998
26.1812 | -35.7435 | -36.1047
26.1873 | -35.7438 | -36.1097
26.1934 | -35.7443 | -36.1143
26.1995 | -35.7446 | -36.1195
26.2056 | -35.7449 | -36.1244
26.2117 | -35.7453 | -36.1294
26.2178 | -35.7457 | -36.1344
26.2239 | -35.746 | -36.1394
26.23 | -35.7463 | -36.1446
26.2361 | -35.7467 | -36.1491
26.2422 | -35.747 | -36.1538
26.2483 | -35.7473 | -36.159
26.2544 | -35.7478 | -36.1635
26.2605 | -35.7481 | -36.1686
26.2666 | -35.7485 | -36.1735
26.2727 | -35.7488 | -36.1783
26.2788 | -35.7491 | -36.1829
26.2849 | -35.7494 | -36.1875
26.2911 | -35.7498 | -36.192
26.2972 | -35.7502 | -36.1967
26.3033 | -35.7506 | -36.2019
26.3094 | -35.7509 | -36.2073
26.3155 | -35.7512 | -36.2114
26.3216 | -35.7515 | -36.2165
26.3277 | -35.7518 | -36.2213
26.3338 | -35.7522 | -36.2261
26.3399 | -35.7527 | -36.2305
26.346 | -35.753 | -36.2345
26.3521 | -35.7534 | -36.2396
26.3582 | -35.7538 | -36.2443
26.3643 | -35.7541 | -36.2493
26.3704 | -35.7545 | -36.2541
26.3765 | -35.7549 | -36.259
26.3826 | -35.7552 | -36.2627
26.3887 | -35.7555 | -36.2668
26.3948 | -35.7558 | -36.2715
26.4009 | -35.7562 | -36.2766
26.407 | -35.7566 | -36.2809
26.4131 | -35.7569 | -36.2858
26.4192 | -35.7572 | -36.2902
26.4253 | -35.7575 | -36.2952
26.4314 | -35.7579 | -36.3004
26.4376 | -35.7582 | -36.3054
26.4437 | -35.7585 | -36.3103
26.4498 | -35.7589 | -36.3153
26.4559 | -35.7591 | -36.3195
26.462 | -35.7595 | -36.3244
26.4681 | -35.7599 | -36.3289
26.4742 | -35.7602 | -36.3338
26.4803 | -35.7605 | -36.3388
26.4864 | -35.7608 | -36.3426
26.4925 | -35.7612 | -36.3474
26.4986 | -35.7615 | -36.3521
26.5047 | -35.7618 | -36.3567
26.5108 | -35.7621 | -36.3613
26.5169 | -35.7625 | -36.3659
26.523 | -35.7628 | -36.3707
26.5291 | -35.7632 | -36.3757
26.5352 | -35.7636 | -36.3802
26.5413 | -35.7639 | -36.3849
26.5474 | -35.7642 | -36.3897
26.5535 | -35.7645 | -36.3946
26.5596 | -35.7647 | -36.3992
26.5657 | -35.7651 | -36.404
26.5718 | -35.7653 | -36.4086
26.578 | -35.7657 | -36.413
26.5841 | -35.7661 | -36.4178
26.5902 | -35.7664 | -36.4227
26.5963 | -35.7666 | -36.4271
26.6024 | -35.767 | -36.4316
26.6085 | -35.7673 | -36.4366
26.6146 | -35.7677 | -36.4416
26.6207 | -35.768 | -36.4467
26.6268 | -35.7683 | -36.4512
26.6329 | -35.7685 | -36.4558
26.639 | -35.7688 | -36.4601
26.6451 | -35.7691 | -36.465
26.6512 | -35.7694 | -36.4703
26.6573 | -35.7697 | -36.4751
26.6634 | -35.77 | -36.48
26.6695 | -35.7703 | -36.4848
26.6756 | -35.7707 | -36.4896
26.6817 | -35.771 | -36.4943
26.6878 | -35.7713 | -36.4993
26.6939 | -35.7717 | -36.5039
26.7 | -35.7721 | -36.5088
26.7061 | -35.7724 | -36.5133
26.7122 | -35.7728 | -36.5182
26.7183 | -35.7731 | -36.5229
26.7245 | -35.7734 | -36.5272
26.7306 | -35.7738 | -36.5323
26.7367 | -35.774 | -36.5364
26.7428 | -35.7744 | -36.5411
26.7489 | -35.7748 | -36.546
26.755 | -35.7751 | -36.5509
26.7611 | -35.7755 | -36.556
26.7672 | -35.7759 | -36.5613
26.7733 | -35.7762 | -36.5662
26.7794 | -35.7765 | -36.5711
26.7855 | -35.7768 | -36.5754
26.7916 | -35.7771 | -36.5798
26.7977 | -35.7774 | -36.5843
26.8038 | -35.7777 | -36.5888
26.8099 | -35.7781 | -36.5935
26.816 | -35.7784 | -36.5983
26.8221 | -35.7787 | -36.6029
26.8282 | -35.779 | -36.6076
26.8343 | -35.7793 | -36.6125
26.8404 | -35.7796 | -36.6171
26.8465 | -35.7799 | -36.6218
26.8526 | -35.7802 | -36.6267
26.8587 | -35.7804 | -36.6314
26.8649 | -35.7808 | -36.6358
26.871 | -35.7811 | -36.6408
26.8771 | -35.7814 | -36.6455
26.8832 | -35.7817 | -36.6501
26.8893 | -35.782 | -36.655
26.8954 | -35.7824 | -36.6596
26.9015 | -35.7826 | -36.6643
26.9076 | -35.7829 | -36.6688
26.9137 | -35.7833 | -36.6735
26.9198 | -35.7836 | -36.6783
26.9259 | -35.7839 | -36.6828
26.932 | -35.7842 | -36.6875
26.9381 | -35.7845 | -36.6919
26.9442 | -35.7848 | -36.6967
26.9503 | -35.7852 | -36.7021
26.9564 | -35.7855 | -36.7073
26.9625 | -35.7859 | -36.7121
26.9686 | -35.7862 | -36.7167
26.9747 | -35.7865 | -36.7216
26.9808 | -35.7867 | -36.7264
26.9869 | -35.787 | -36.7317
26.993 | -35.7873 | -36.7368
26.9991 | -35.7877 | -36.7414
27.0052 | -35.788 | -36.7459
27.0114 | -35.7882 | -36.7505
27.0175 | -35.7885 | -36.7556
27.0236 | -35.7888 | -36.7602
27.0297 | -35.7891 | -36.765
27.0358 | -35.7894 | -36.7696
27.0419 | -35.7898 | -36.7742
27.048 | -35.7901 | -36.7788
27.0541 | -35.7904 | -36.7838
27.0602 | -35.7906 | -36.7884
27.0663 | -35.7909 | -36.7932
27.0724 | -35.7912 | -36.7982
27.0785 | -35.7915 | -36.803
27.0846 | -35.7917 | -36.8076
27.0907 | -35.7921 | -36.8124
27.0968 | -35.7924 | -36.817
27.1029 | -35.7927 | -36.822
27.109 | -35.793 | -36.8267
27.1151 | -35.7933 | -36.8312
27.1212 | -35.7936 | -36.8363
27.1273 | -35.7939 | -36.8407
27.1334 | -35.7942 | -36.8454
27.1395 | -35.7945 | -36.85
27.1456 | -35.7949 | -36.8548
27.1518 | -35.7952 | -36.8594
27.1579 | -35.7954 | -36.8635
27.164 | -35.7957 | -36.868
27.1701 | -35.796 | -36.8727
27.1762 | -35.7964 | -36.8771
27.1823 | -35.7966 | -36.8813
27.1884 | -35.7969 | -36.8863
27.1945 | -35.7972 | -36.8909
27.2006 | -35.7974 | -36.8953
27.2067 | -35.7978 | -36.9002
27.2128 | -35.7981 | -36.9048
27.2189 | -35.7983 | -36.9095
27.225 | -35.7986 | -36.9145
27.2311 | -35.7989 | -36.9192
27.2372 | -35.7992 | -36.9238
27.2433 | -35.7995 | -36.9287
27.2494 | -35.7999 | -36.9332
27.2555 | -35.8002 | -36.9377
27.2616 | -35.8005 | -36.942
27.2677 | -35.8008 | -36.9467
27.2738 | -35.8011 | -36.9512
27.2799 | -35.8015 | -36.9555
27.286 | -35.8017 | -36.96
27.2921 | -35.802 | -36.9646
27.2983 | -35.8022 | -36.9695
27.3044 | -35.8025 | -36.974
27.3105 | -35.8028 | -36.9785
27.3166 | -35.8031 | -36.9834
27.3227 | -35.8034 | -36.988
27.3288 | -35.8036 | -36.9928
27.3349 | -35.8039 | -36.9969
27.341 | -35.8042 | -37.0013
27.3471 | -35.8045 | -37.0059
27.3532 | -35.8048 | -37.0104
27.3593 | -35.8051 | -37.0152
27.3654 | -35.8054 | -37.02
27.3715 | -35.8057 | -37.0236
27.3776 | -35.8059 | -37.0281
27.3837 | -35.8062 | -37.0327
27.3898 | -35.8065 | -37.0374
27.3959 | -35.8067 | -37.0421
27.402 | -35.807 | -37.0464
27.4081 | -35.8073 | -37.0509
27.4142 | -35.8077 | -37.0555
27.4203 | -35.808 | -37.0602
27.4264 | -35.8083 | -37.0647
27.4325 | -35.8086 | -37.0691
27.4387 | -35.8089 | -37.074
27.4448 | -35.8091 | -37.0785
27.4509 | -35.8094 | -37.0833
27.457 | -35.8098 | -37.0879
27.4631 | -35.81 | -37.0926
27.4692 | -35.8103 | -37.0969
27.4753 | -35.8104 | -37.1015
27.4814 | -35.8107 | -37.1062
27.4875 | -35.811 | -37.1111
27.4936 | -35.8113 | -37.1161
27.4997 | -35.8117 | -37.1212
27.5058 | -35.812 | -37.1259
27.5119 | -35.8123 | -37.1304
27.518 | -35.8125 | -37.1348
27.5241 | -35.8128 | -37.1394
27.5302 | -35.813 | -37.1444
27.5363 | -35.8133 | -37.1492
27.5424 | -35.8135 | -37.1536
27.5485 | -35.8138 | -37.158
27.5546 | -35.8141 | -37.1622
27.5607 | -35.8144 | -37.1666
27.5668 | -35.8147 | -37.171
27.5729 | -35.8149 | -37.1755
27.5791 | -35.8152 | -37.1803
27.5852 | -35.8154 | -37.1852
27.5913 | -35.8157 | -37.1895
27.5974 | -35.816 | -37.1941
27.6035 | -35.8163 | -37.1985
27.6096 | -35.8166 | -37.2029
27.6157 | -35.8169 | -37.2071
27.6218 | -35.8172 | -37.212
27.6279 | -35.8174 | -37.2167
27.634 | -35.8177 | -37.2214
27.6401 | -35.8179 | -37.2263
27.6462 | -35.8183 | -37.2308
27.6523 | -35.8185 | -37.2352
27.6584 | -35.8187 | -37.2394
27.6645 | -35.819 | -37.2438
27.6706 | -35.8193 | -37.2491
27.6767 | -35.8195 | -37.2534
27.6828 | -35.8197 | -37.2576
27.6889 | -35.8201 | -37.2625
27.695 | -35.8204 | -37.267
27.7011 | -35.8206 | -37.2717
27.7072 | -35.8209 | -37.2766
27.7133 | -35.8211 | -37.2809
27.7194 | -35.8213 | -37.2854
27.7256 | -35.8216 | -37.29
27.7317 | -35.8218 | -37.2946
27.7378 | -35.8222 | -37.2989
27.7439 | -35.8225 | -37.3038
27.75 | -35.8228 | -37.3086
27.7561 | -35.8229 | -37.3131
27.7622 | -35.8232 | -37.318
27.7683 | -35.8235 | -37.3226
27.7744 | -35.8238 | -37.3273
27.7805 | -35.8241 | -37.3318
27.7866 | -35.8243 | -37.3363
27.7927 | -35.8247 | -37.3415
27.7988 | -35.8249 | -37.3461
27.8049 | -35.8252 | -37.3503
27.811 | -35.8255 | -37.3551
27.8171 | -35.8257 | -37.3594
27.8232 | -35.826 | -37.3639
27.8293 | -35.8262 | -37.3681
27.8354 | -35.8265 | -37.3725
27.8415 | -35.8268 | -37.3774
27.8476 | -35.827 | -37.382
27.8537 | -35.8272 | -37.3863
27.8598 | -35.8275 | -37.3906
27.866 | -35.8278 | -37.3952
27.8721 | -35.828 | -37.4
27.8782 | -35.8283 | -37.4046
27.8843 | -35.8285 | -37.4096
27.8904 | -35.8288 | -37.4145
27.8965 | -35.829 | -37.4192
27.9026 | -35.8294 | -37.4244
27.9087 | -35.8297 | -37.4295
27.9148 | -35.8299 | -37.4343
27.9209 | -35.8302 | -37.4385
27.927 | -35.8306 | -37.4436
27.9331 | -35.8308 | -37.4485
27.9392 | -35.8311 | -37.4534
27.9453 | -35.8314 | -37.458
27.9514 | -35.8316 | -37.4629
27.9575 | -35.8319 | -37.4682
27.9636 | -35.8321 | -37.4731
27.9697 | -35.8323 | -37.4779
27.9758 | -35.8326 | -37.4824
27.9819 | -35.8328 | -37.4874
27.988 | -35.8332 | -37.492
27.9941 | -35.8334 | -37.4966
28.0002 | -35.8337 | -37.5008
28.0063 | -35.834 | -37.505
28.0125 | -35.8342 | -37.509
28.0186 | -35.8345 | -37.5139
28.0247 | -35.8347 | -37.5183
28.0308 | -35.8349 | -37.5232
28.0369 | -35.8352 | -37.5288
28.043 | -35.8355 | -37.5333
28.0491 | -35.8357 | -37.5382
28.0552 | -35.8361 | -37.5431
28.0613 | -35.8363 | -37.5475
28.0674 | -35.8365 | -37.5518
28.0735 | -35.8367 | -37.557

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.3678 | -41.2434
21.4908 | -35.5607 | -34.6215
16.4948 | -34.9457 | -27.3155
12.6602 | -34.5914 | -21.108
9.7171 | -34.3785 | -16.2583
7.4582 | -34.2282 | -12.3243
5.7244 | -34.1229 | -9.1408
4.3936 | -34.0673 | -6.5641
3.3722 | -34.0523 | -4.54
2.5883 | -34.0529 | -3.0038
1.9866 | -34.0488 | -1.7761
1.5248 | -34.06 | -0.79353
1.1703 | -34.0673 | -0.15318
1 | -34.0661 | 0.16523
0.89822 | -34.0588 | 0.4325
0.68941 | -34.0602 | 1.0095
0.52919 | -34.0465 | 1.4539
0.40612 | -34.0639 | 2.055
0.31171 | -34.0517 | 2.5024
0.23928 | -34.0432 | 3.2333
0.18358 | -34.0638 | 3.6414
0.14091 | -34.067 | 4.5275
0.10819 | -34.0922 | 5.6817
0.083017 | -34.1162 | 7.1542
0.063705 | -34.1739 | 9.0117
0.048923 | -34.2318 | 11.4913
0.037479 | -34.3415 | 14.6721
0.028801 | -34.5046 | 18.6921
0.022078 | -34.8251 | 24.1295
0.017023 | -35.308 | 30.5615
0.013018 | -36.1124 | 38.2297
0.010014 | -37.2279 | 46.1623

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003673 | -2.8394
0.007347 | -26.5637
0.01102 | -35.2868
0.014693 | -40.8107
0.018366 | -43.4315
0.02204 | -45.3377
0.025713 | -46.1823
0.029386 | -48.6408
0.033059 | -49.7421
0.036733 | -50.0751
0.040406 | -52.644
0.044079 | -54.4197
0.047752 | -56.1439
0.051426 | -55.7148
0.055099 | -55.7436
0.058772 | -56.4151
0.062445 | -57.5076
0.066119 | -59.9795
0.069792 | -60.5589
0.073465 | -59.3848
0.077139 | -59.8188
0.080812 | -60.9094
0.084485 | -60.8839
0.088158 | -61.049
0.091832 | -61.7535
0.095505 | -63.6106
0.099178 | -64.5591
0.10285 | -64.7358
0.10652 | -64.6318
0.1102 | -64.7798
0.11387 | -65.8676
0.11754 | -67.201
0.12122 | -67.2836
0.12489 | -66.7925
0.12856 | -68.4602
0.13224 | -68.1199
0.13591 | -67.1315
0.13958 | -68.0658
0.14326 | -68.5226
0.14693 | -68.9777
0.1506 | -68.9915
0.15428 | -69.7045
0.15795 | -70.2309
0.16162 | -69.3366
0.1653 | -69.0777
0.16897 | -70.394
0.17264 | -70.0462
0.17632 | -69.7945
0.17999 | -70.7714
0.18366 | -70.7654
0.18734 | -71.4374
0.19101 | -71.0351
0.19468 | -70.5441
0.19836 | -70.6312
0.20203 | -70.9717
0.2057 | -71.4599
0.20938 | -71.7969
0.21305 | -72.3483
0.21672 | -72.2711
0.2204 | -72.7973
0.22407 | -72.6649
0.22774 | -72.2483
0.23142 | -72.3402
0.23509 | -73.8523
0.23876 | -74.5502
0.24244 | -73.7732
0.24611 | -74.3181
0.24978 | -74.4673
0.25345 | -74.1985
0.25713 | -74.914
0.2608 | -75.3252
0.26448 | -74.8784
0.26815 | -74.149
0.27182 | -76.008
0.27549 | -77.1626
0.27917 | -75.4644
0.28284 | -75.705
0.28652 | -76.4405
0.29019 | -74.8425
0.29386 | -75.6357
0.29753 | -76.2106
0.30121 | -75.8086
0.30488 | -75.5525
0.30855 | -76.114
0.31223 | -76.0705
0.3159 | -76.7409
0.31957 | -78.8597
0.32325 | -77.9677
0.32692 | -77.1242
0.33059 | -76.8875
0.33427 | -75.981
0.33794 | -76.3748
0.34161 | -77.7309
0.34529 | -77.4998
0.34896 | -77.9229
0.35263 | -78.1304
0.35631 | -77.1308
0.35998 | -77.3833
0.36365 | -77.9904
0.36733 | -77.591
0.371 | -78.211
0.37467 | -78.7051
0.37835 | -78.1996
0.38202 | -79.3097
0.38569 | -79.6802
0.38937 | -80.0989
0.39304 | -79.8217
0.39671 | -78.5572
0.40039 | -78.952
0.40406 | -78.879
0.40773 | -78.9717
0.41141 | -79.3548
0.41508 | -79.03
0.41875 | -78.2428
0.42242 | -78.8866
0.4261 | -80.3512
0.42977 | -80.5428
0.43345 | -79.7295
0.43712 | -80.04
0.44079 | -81.934
0.44446 | -81.8687
0.44814 | -80.6503
0.45181 | -80.1179
0.45548 | -80.5079
0.45916 | -80.1162
0.46283 | -80.0128
0.4665 | -80.7343
0.47018 | -81.2775
0.47385 | -80.8687
0.47752 | -80.0414
0.4812 | -80.2952
0.48487 | -80.7002
0.48854 | -81.434
0.49222 | -81.3793
0.49589 | -81.8326
0.49956 | -80.5291
0.50324 | -80.5564
0.50691 | -80.8771
0.51058 | -81.7653
0.51426 | -82.7945
0.51793 | -82.2136
0.5216 | -81.8281
0.52528 | -80.9116
0.52895 | -81.1869
0.53262 | -81.1237
0.5363 | -80.8878
0.53997 | -81.2166
0.54364 | -81.7189
0.54732 | -81.7135
0.55099 | -81.6611
0.55466 | -82.1627
0.55834 | -82.4382
0.56201 | -82.1013
0.56568 | -81.917
0.56936 | -81.9859
0.57303 | -82.2264
0.5767 | -82.3854
0.58038 | -82.5952
0.58405 | -82.937
0.58772 | -83.3798
0.5914 | -83.3979
0.59507 | -83.1223
0.59874 | -83.0533
0.60242 | -83.3677
0.60609 | -83.3336
0.60976 | -83.2618
0.61343 | -83.3361
0.61711 | -83.3596
0.62078 | -83.1501
0.62445 | -82.958
0.62813 | -83.0524
0.6318 | -83.1011
0.63548 | -83.0467
0.63915 | -83.1399
0.64282 | -83.0417
0.64649 | -82.9022
0.65017 | -82.7775
0.65384 | -83.0742
0.65751 | -83.5457
0.66119 | -83.4791
0.66486 | -83.4693
0.66853 | -84.1264
0.67221 | -84.723
0.67588 | -85.1181
0.67955 | -85.4575
0.68323 | -85.0658
0.6869 | -84.5794
0.69057 | -84.5116
0.69425 | -84.6784
0.69792 | -84.7535
0.70159 | -84.4731
0.70527 | -84.0617
0.70894 | -84.0664
0.71261 | -84.4149
0.71629 | -84.3317
0.71996 | -84.3566
0.72363 | -84.8551
0.72731 | -85.2701
0.73098 | -85.338
0.73465 | -85.1901
0.73833 | -84.8782
0.742 | -84.8036
0.74567 | -84.7309
0.74935 | -84.5657
0.75302 | -84.9124
0.75669 | -85.1508
0.76036 | -85.0074
0.76404 | -85.3062
0.76771 | -85.82
0.77138 | -86.3516
0.77506 | -86.5967
0.77873 | -85.9989
0.78241 | -85.2664
0.78608 | -85.2236
0.78975 | -85.536
0.79343 | -85.7221
0.7971 | -85.9453
0.80077 | -85.9715
0.80444 | -85.6078
0.80812 | -85.5883
0.81179 | -85.6101
0.81546 | -85.8794
0.81914 | -86.4193
0.82281 | -86.5857
0.82648 | -85.971
0.83016 | -85.6098
0.83383 | -86.0118
0.8375 | -86.3811
0.84118 | -86.2744
0.84485 | -86.1835
0.84852 | -86.1063
0.8522 | -85.8687
0.85587 | -85.8817
0.85954 | -86.3029
0.86322 | -86.6003
0.86689 | -86.6454
0.87056 | -86.3965
0.87424 | -85.9809
0.87791 | -85.71
0.88158 | -85.9462
0.88526 | -86.319
0.88893 | -86.7412
0.8926 | -87.1548
0.89628 | -87.2686
0.89995 | -87.1696
0.90362 | -86.9373
0.9073 | -86.6765
0.91097 | -86.6057
0.91464 | -86.2816
0.91832 | -85.9336
0.92199 | -86.1188
0.92566 | -86.7643
0.92934 | -87.4363
0.93301 | -87.6069
0.93668 | -87.103
0.94036 | -86.4439
0.94403 | -86.3917
0.9477 | -86.7167
0.95137 | -86.93
0.95505 | -86.8316
0.95872 | -86.6591
0.9624 | -86.9514
0.96607 | -87.2064
0.96974 | -87.1307
0.97342 | -87.2648
0.97709 | -87.1668
0.98076 | -86.8646
0.98443 | -86.8761
0.98811 | -86.8875
0.99178 | -86.8989
0.99545 | -86.9103
0.99913 | -86.9217
1.0028 | -86.9331
1.0065 | -86.9445
1.0101 | -86.956
1.0138 | -86.9674
1.0175 | -86.9788
1.0212 | -86.9902
1.0248 | -87.4365
1.0285 | -87.2691
1.0322 | -87.6707
1.0359 | -88.6029
1.0395 | -88.9872
1.0432 | -88.4317
1.0469 | -87.8601
1.0506 | -87.5866
1.0542 | -87.7007
1.0579 | -88.1688
1.0616 | -88.6408
1.0652 | -89.015
1.0689 | -89.1998
1.0726 | -89.2553
1.0763 | -89.1649
1.0799 | -89.0865
1.0836 | -89.2643
1.0873 | -89.5031
1.091 | -89.32
1.0946 | -88.7804
1.0983 | -88.4916
1.102 | -88.6784
1.1057 | -88.983
1.1093 | -89.045
1.113 | -88.8675
1.1167 | -88.5845
1.1203 | -88.4133
1.124 | -88.5076
1.1277 | -88.8194
1.1314 | -89.2394
1.135 | -89.5772
1.1387 | -89.5923
1.1424 | -89.3605
1.1461 | -89.052
1.1497 | -88.7635
1.1534 | -88.6772
1.1571 | -88.8777
1.1608 | -89.2261
1.1644 | -89.422
1.1681 | -89.3354
1.1718 | -89.2942
1.1754 | -89.4984
1.1791 | -89.6334
1.1828 | -89.3182
1.1865 | -88.7111
1.1901 | -88.3209
1.1938 | -88.4231
1.1975 | -88.8386
1.2012 | -89.1856
1.2048 | -89.27
1.2085 | -89.1813
1.2122 | -89.1324
1.2159 | -89.2378
1.2195 | -89.3777
1.2232 | -89.4205
1.2269 | -89.4098
1.2305 | -89.4198
1.2342 | -89.4035
1.2379 | -89.3693
1.2416 | -89.4186
1.2452 | -89.4466
1.2489 | -89.3527
1.2526 | -89.4019
1.2563 | -89.687
1.2599 | -89.9245
1.2636 | -89.9511
1.2673 | -89.8512
1.2709 | -89.742
1.2746 | -89.6562
1.2783 | -89.6398
1.282 | -89.8441
1.2856 | -90.1458
1.2893 | -90.2905
1.293 | -90.3776
1.2967 | -90.4724
1.3003 | -90.5051
1.304 | -90.5388
1.3077 | -90.5827
1.3114 | -90.5835
1.315 | -90.4794
1.3187 | -90.2405
1.3224 | -89.994
1.326 | -89.8416
1.3297 | -89.7876
1.3334 | -89.8926
1.3371 | -90.155
1.3407 | -90.4051
1.3444 | -90.5615
1.3481 | -90.7579
1.3518 | -90.9606
1.3554 | -90.9546
1.3591 | -90.7774
1.3628 | -90.6208
1.3665 | -90.5964
1.3701 | -90.6672
1.3738 | -90.6377
1.3775 | -90.4137
1.3811 | -90.2088
1.3848 | -90.2221
1.3885 | -90.3844
1.3922 | -90.5302
1.3958 | -90.5793
1.3995 | -90.5548
1.4032 | -90.6234
1.4069 | -90.8265
1.4105 | -90.975
1.4142 | -90.9787
1.4179 | -90.8731
1.4216 | -90.6882
1.4252 | -90.6117
1.4289 | -90.8139
1.4326 | -91.1479
1.4362 | -91.3842
1.4399 | -91.5366
1.4436 | -91.6144
1.4473 | -91.4936
1.4509 | -91.2068
1.4546 | -90.9111
1.4583 | -90.6636
1.462 | -90.4825
1.4656 | -90.5165
1.4693 | -90.8168
1.473 | -91.1583
1.4767 | -91.3549
1.4803 | -91.4866
1.484 | -91.6465
1.4877 | -91.7711
1.4913 | -91.8382
1.495 | -91.9877
1.4987 | -92.1245
1.5024 | -91.999
1.506 | -91.7439
1.5097 | -91.592
1.5134 | -91.5954
1.5171 | -91.6357
1.5207 | -91.4877
1.5244 | -91.227
1.5281 | -91.1206
1.5318 | -91.1546
1.5354 | -91.2004
1.5391 | -91.2375
1.5428 | -91.2271
1.5464 | -91.222
1.5501 | -91.3461
1.5538 | -91.4252
1.5575 | -91.2867
1.5611 | -91.1058
1.5648 | -91.0439
1.5685 | -91.1627
1.5722 | -91.4815
1.5758 | -91.8744
1.5795 | -92.1247
1.5832 | -92.2117
1.5869 | -92.2904
1.5905 | -92.3551
1.5942 | -92.3969
1.5979 | -92.4419
1.6015 | -92.523
1.6052 | -92.6062
1.6089 | -92.615
1.6126 | -92.512
1.6162 | -92.3585
1.6199 | -92.2121
1.6236 | -92.0893
1.6273 | -91.9864
1.6309 | -91.9316
1.6346 | -91.9493
1.6383 | -92.0253
1.6419 | -92.1042
1.6456 | -92.1595
1.6493 | -92.2054
1.653 | -92.2815
1.6566 | -92.3832
1.6603 | -92.426
1.664 | -92.3102
1.6677 | -92.1236
1.6713 | -92.0628
1.675 | -92.1865
1.6787 | -92.3511
1.6824 | -92.3522
1.686 | -92.1506
1.6897 | -91.9159
1.6934 | -91.8593
1.697 | -92.042
1.7007 | -92.356
1.7044 | -92.6259
1.7081 | -92.7162
1.7117 | -92.5829
1.7154 | -92.375
1.7191 | -92.3215
1.7228 | -92.5108
1.7264 | -92.8065
1.7301 | -92.9919
1.7338 | -92.9981
1.7375 | -92.9323
1.7411 | -92.8892
1.7448 | -92.8559
1.7485 | -92.7676
1.7521 | -92.6652
1.7558 | -92.6956
1.7595 | -92.9037
1.7632 | -93.1085
1.7668 | -93.1029
1.7705 | -92.8758
1.7742 | -92.5781
1.7779 | -92.3187
1.7815 | -92.1583
1.7852 | -92.1466
1.7889 | -92.2685
1.7926 | -92.4461
1.7962 | -92.6066
1.7999 | -92.6968
1.8036 | -92.7162
1.8072 | -92.6902
1.8109 | -92.6882
1.8146 | -92.7272
1.8183 | -92.7296
1.8219 | -92.6115
1.8256 | -92.4035
1.8293 | -92.2509
1.833 | -92.2938
1.8366 | -92.5773
1.8403 | -93.0216
1.844 | -93.4186
1.8477 | -93.6055
1.8513 | -93.603
1.855 | -93.5397
1.8587 | -93.4984
1.8623 | -93.4445
1.866 | -93.3517
1.8697 | -93.2569
1.8734 | -93.2183
1.877 | -93.2465
1.8807 | -93.2868
1.8844 | -93.2841
1.8881 | -93.2396
1.8917 | -93.191
1.8954 | -93.1714
1.8991 | -93.1583
1.9027 | -93.1373
1.9064 | -93.0734
1.9101 | -92.9274
1.9138 | -92.7815
1.9174 | -92.7773
1.9211 | -93.0023
1.9248 | -93.446
1.9285 | -93.972
1.9321 | -94.3773
1.9358 | -94.5382
1.9395 | -94.5358
1.9432 | -94.5228
1.9468 | -94.5199
1.9505 | -94.458
1.9542 | -94.3049
1.9578 | -94.1295
1.9615 | -93.9963
1.9652 | -93.8694
1.9689 | -93.7178
1.9725 | -93.6094
1.9762 | -93.6182
1.9799 | -93.7498
1.9836 | -93.927
1.9872 | -94.1026
1.9909 | -94.2983
1.9946 | -94.4616
1.9983 | -94.3926
2.0019 | -94.0279
2.0056 | -93.6005
2.0093 | -93.3533
2.0129 | -93.3043
2.0166 | -93.339
2.0203 | -93.3702
2.024 | -93.4139
2.0276 | -93.5313
2.0313 | -93.7299
2.035 | -93.9421
2.0387 | -94.0956
2.0423 | -94.161
2.046 | -94.1051
2.0497 | -93.9556
2.0534 | -93.7818
2.057 | -93.6202
2.0607 | -93.4833
2.0644 | -93.4208
2.068 | -93.4924
2.0717 | -93.6694
2.0754 | -93.8711
2.0791 | -94.0869
2.0827 | -94.3004
2.0864 | -94.4197
2.0901 | -94.3348
2.0938 | -94.0883
2.0974 | -93.8615
2.1011 | -93.7757
2.1048 | -93.8427
2.1085 | -93.9377
2.1121 | -93.9539
2.1158 | -93.9111
2.1195 | -93.9681
2.1231 | -94.0386
2.1268 | -94.1229
2.1305 | -94.16
2.1342 | -94.1303
2.1378 | -94.0782
2.1415 | -94.0611
2.1452 | -94.1221
2.1489 | -94.2191
2.1525 | -94.3077
2.1562 | -94.3655
2.1599 | -94.3731
2.1636 | -94.3548
2.1672 | -94.3394
2.1709 | -94.3564
2.1746 | -94.3834
2.1782 | -94.3821
2.1819 | -94.3299
2.1856 | -94.2313
2.1893 | -94.1321
2.1929 | -94.0571
2.1966 | -94.0137
2.2003 | -94.0196
2.204 | -94.0607
2.2076 | -94.1367
2.2113 | -94.2266
2.215 | -94.3145
2.2187 | -94.4023
2.2223 | -94.514
2.226 | -94.667
2.2297 | -94.8264
2.2333 | -94.9462
2.237 | -94.9618
2.2407 | -94.8616
2.2444 | -94.7132
2.248 | -94.5739
2.2517 | -94.4939
2.2554 | -94.4895
2.2591 | -94.5543
2.2627 | -94.6623
2.2664 | -94.7973
2.2701 | -94.9447
2.2738 | -95.0898
2.2774 | -95.2078
2.2811 | -95.2771
2.2848 | -95.2992
2.2884 | -95.3059
2.2921 | -95.3078
2.2958 | -95.2707
2.2995 | -95.1672
2.3031 | -95.0105
2.3068 | -94.8422
2.3105 | -94.7173
2.3142 | -94.6585
2.3178 | -94.681
2.3215 | -94.7708
2.3252 | -94.8973
2.3288 | -95.0183
2.3325 | -95.1061
2.3362 | -95.1553
2.3399 | -95.1397
2.3435 | -95.033
2.3472 | -94.8723
2.3509 | -94.7375
2.3546 | -94.7294
2.3582 | -94.8758
2.3619 | -95.1256
2.3656 | -95.3761
2.3693 | -95.5329
2.3729 | -95.5654
2.3766 | -95.5196
2.3803 | -95.4769
2.3839 | -95.4877
2.3876 | -95.5256
2.3913 | -95.551
2.395 | -95.5063
2.3986 | -95.4067
2.4023 | -95.3033
2.406 | -95.2501
2.4097 | -95.259
2.4133 | -95.3093
2.417 | -95.3441
2.4207 | -95.3401
2.4244 | -95.3014
2.428 | -95.2471
2.4317 | -95.1686
2.4354 | -95.055
2.439 | -94.9235
2.4427 | -94.8263
2.4464 | -94.7914
2.4501 | -94.8318
2.4537 | -94.9314
2.4574 | -95.0924
2.4611 | -95.2778
2.4648 | -95.4553
2.4684 | -95.6036
2.4721 | -95.7256
2.4758 | -95.8226
2.4795 | -95.8971
2.4831 | -95.9498
2.4868 | -95.9641
2.4905 | -95.9063
2.4941 | -95.7497
2.4978 | -95.5058
2.5015 | -95.2589
2.5052 | -95.1035
2.5088 | -95.0924
2.5125 | -95.2056
2.5162 | -95.3652
2.5199 | -95.5016
2.5235 | -95.5667
2.5272 | -95.5658
2.5309 | -95.504
2.5346 | -95.4396
2.5382 | -95.4163
2.5419 | -95.4547
2.5456 | -95.5237
2.5492 | -95.588
2.5529 | -95.6387
2.5566 | -95.6568
2.5603 | -95.6435
2.5639 | -95.618
2.5676 | -95.5734
2.5713 | -95.5212
2.575 | -95.4682
2.5786 | -95.4409
2.5823 | -95.4334
2.586 | -95.4369
2.5897 | -95.4365
2.5933 | -95.4116
2.597 | -95.3774
2.6007 | -95.3453
2.6043 | -95.3542
2.608 | -95.411
2.6117 | -95.4895
2.6154 | -95.5613
2.619 | -95.607
2.6227 | -95.6412
2.6264 | -95.6959
2.6301 | -95.8068
2.6337 | -95.9963
2.6374 | -96.2204
2.6411 | -96.3982
2.6448 | -96.4514
2.6484 | -96.2535
2.6521 | -96.1075
2.6558 | -95.9462
2.6594 | -95.823
2.6631 | -95.7773
2.6668 | -95.8346
2.6705 | -95.9779
2.6741 | -96.1564
2.6778 | -96.2964
2.6815 | -96.3545
2.6852 | -96.3319
2.6888 | -96.2524
2.6925 | -96.14
2.6962 | -96.0295
2.6998 | -95.9159
2.7035 | -95.8168
2.7072 | -95.7642
2.7109 | -95.7831
2.7145 | -95.8677
2.7182 | -95.9972
2.7219 | -96.1342
2.7256 | -96.2428
2.7292 | -96.2943
2.7329 | -96.2578
2.7366 | -96.1465
2.7403 | -96.0024
2.7439 | -95.8779
2.7476 | -95.8124
2.7513 | -95.8318
2.7549 | -95.9063
2.7586 | -95.983
2.7623 | -96.0241
2.766 | -96.0232
2.7696 | -96.0111
2.7733 | -96.0137
2.777 | -96.0386
2.7807 | -96.0901
2.7843 | -96.1357
2.788 | -96.148
2.7917 | -96.1117
2.7954 | -96.0227
2.799 | -95.8847
2.8027 | -95.7368
2.8064 | -95.6405
2.81 | -95.6322
2.8137 | -95.724
2.8174 | -95.8888
2.8211 | -96.0755
2.8247 | -96.2102
2.8284 | -96.2166
2.8321 | -96.0902
2.8358 | -95.8693
2.8394 | -95.6519
2.8431 | -95.5349
2.8468 | -95.5719
2.8505 | -95.7747
2.8541 | -96.0921
2.8578 | -96.4523
2.8615 | -96.7674
2.8651 | -96.9892
2.8688 | -97.088
2.8725 | -97.0741
2.8762 | -96.9847
2.8798 | -96.8577
2.8835 | -96.7412
2.8872 | -96.6349
2.8909 | -96.5454
2.8945 | -96.465
2.8982 | -96.3894
2.9019 | -96.3452
2.9056 | -96.3316
2.9092 | -96.346
2.9129 | -96.3824
2.9166 | -96.4488
2.9202 | -96.5371
2.9239 | -96.6324
2.9276 | -96.7126
2.9313 | -96.7823
2.9349 | -96.8401
2.9386 | -96.879
2.9423 | -96.8832
2.946 | -96.8487
2.9496 | -96.7443
2.9533 | -96.5821
2.957 | -96.4036
2.9606 | -96.2441
2.9643 | -96.1517
2.968 | -96.1493
2.9717 | -96.2047
2.9753 | -96.2147
2.979 | -96.069
2.9827 | -96.069
2.9864 | -96.1429
2.99 | -96.2168
2.9937 | -96.2908
2.9974 | -96.3647
3.0011 | -96.4387
3.0047 | -96.5126
3.0084 | -96.5866
3.0121 | -96.6605
3.0157 | -96.7345
3.0194 | -96.8084
3.0231 | -97.0725
3.0268 | -97.394
3.0304 | -97.4969
3.0341 | -97.4987
3.0378 | -97.4279
3.0415 | -97.338
3.0451 | -97.2722
3.0488 | -97.2636
3.0525 | -97.2967
3.0562 | -97.3375
3.0598 | -97.3682
3.0635 | -97.3557
3.0672 | -97.302
3.0708 | -97.2186
3.0745 | -97.1003
3.0782 | -96.9772
3.0819 | -96.8661
3.0855 | -96.784
3.0892 | -96.7179
3.0929 | -96.6673
3.0966 | -96.6236
3.1002 | -96.5885
3.1039 | -96.571
3.1076 | -96.5625
3.1113 | -96.5607
3.1149 | -96.5386
3.1186 | -96.4861
3.1223 | -96.4187
3.1259 | -96.356
3.1296 | -96.331
3.1333 | -96.3637
3.137 | -96.4563
3.1406 | -96.5917
3.1443 | -96.7554
3.148 | -96.9113
3.1517 | -97.0177
3.1553 | -97.0657
3.159 | -97.091
3.1627 | -97.1284
3.1664 | -97.194
3.17 | -97.2867
3.1737 | -97.3911
3.1774 | -97.4583
3.181 | -97.517
3.1847 | -97.5528
3.1884 | -97.5681
3.1921 | -97.5734
3.1957 | -97.5862
3.1994 | -97.62
3.2031 | -97.6904
3.2068 | -97.7852
3.2104 | -97.8776
3.2141 | -97.958
3.2178 | -98.0153
3.2215 | -98.0436
3.2251 | -98.0504
3.2288 | -98.0257
3.2325 | -97.9792
3.2361 | -97.9252
3.2398 | -97.8665
3.2435 | -97.812
3.2472 | -97.7609
3.2508 | -97.7248
3.2545 | -97.7107
3.2582 | -97.7034
3.2619 | -97.7045
3.2655 | -97.6911
3.2692 | -97.6446
3.2729 | -97.5636
3.2766 | -97.4471
3.2802 | -97.3212
3.2839 | -97.2193
3.2876 | -97.1533
3.2912 | -97.1302
3.2949 | -97.149
3.2986 | -97.1968
3.3023 | -97.2687
3.3059 | -97.3502
3.3096 | -97.4246
3.3133 | -97.4775
3.317 | -97.5044
3.3206 | -97.5161
3.3243 | -97.5202
3.328 | -97.5211
3.3316 | -97.5376
3.3353 | -97.572
3.339 | -97.6246
3.3427 | -97.6898
3.3463 | -97.7625
3.35 | -97.8326
3.3537 | -97.8934
3.3574 | -97.9324
3.361 | -97.96
3.3647 | -97.9865
3.3684 | -98.011
3.3721 | -98.023
3.3757 | -98.0294
3.3794 | -98.0221
3.3831 | -97.9967
3.3867 | -97.9444
3.3904 | -97.8794
3.3941 | -97.8176
3.3978 | -97.768
3.4014 | -97.7285
3.4051 | -97.711
3.4088 | -97.7083
3.4125 | -97.7322
3.4161 | -97.7805
3.4198 | -97.866
3.4235 | -97.9803
3.4272 | -98.0918
3.4308 | -98.176
3.4345 | -98.2253
3.4382 | -98.2394
3.4418 | -98.2143
3.4455 | -98.1555
3.4492 | -98.0865
3.4529 | -98.02
3.4565 | -97.965
3.4602 | -97.9191
3.4639 | -97.8968
3.4676 | -97.8998
3.4712 | -97.9186
3.4749 | -97.9383
3.4786 | -97.951
3.4823 | -97.9502
3.4859 | -97.9104
3.4896 | -97.8413
3.4933 | -97.7551
3.4969 | -97.6859
3.5006 | -97.6614
3.5043 | -97.6974
3.508 | -97.8063
3.5116 | -97.9769
3.5153 | -98.1922
3.519 | -98.4094
3.5227 | -98.5949
3.5263 | -98.7335
3.53 | -98.7951
3.5337 | -98.7854
3.5374 | -98.722
3.541 | -98.6158
3.5447 | -98.4796
3.5484 | -98.3442
3.552 | -98.218
3.5557 | -98.1162
3.5594 | -98.048
3.5631 | -98.0144
3.5667 | -98.0203
3.5704 | -98.0644
3.5741 | -98.1487
3.5778 | -98.2727
3.5814 | -98.4168
3.5851 | -98.5712
3.5888 | -98.712
3.5925 | -98.8079
3.5961 | -98.8662
3.5998 | -98.8893
3.6035 | -98.8826
3.6071 | -98.8601
3.6108 | -98.8124
3.6145 | -98.7479
3.6182 | -98.6556
3.6218 | -98.5535
3.6255 | -98.4557
3.6292 | -98.3777
3.6329 | -98.3404
3.6365 | -98.3602
3.6402 | -98.4236
3.6439 | -98.5214
3.6475 | -98.647
3.6512 | -98.7649
3.6549 | -98.8598
3.6586 | -98.921
3.6622 | -98.9385
3.6659 | -98.9103
3.6696 | -98.8496
3.6733 | -98.7742
3.6769 | -98.6852
3.6806 | -98.5997
3.6843 | -98.5198
3.688 | -98.4451
3.6916 | -98.3812
3.6953 | -98.3315
3.699 | -98.2981
3.7026 | -98.2879
3.7063 | -98.3416
3.71 | -98.386
3.7137 | -98.4429
3.7173 | -98.5021
3.721 | -98.5594
3.7247 | -98.616
3.7284 | -98.6687
3.732 | -98.7156
3.7357 | -98.763
3.7394 | -98.8229
3.7431 | -98.8939
3.7467 | -98.973
3.7504 | -99.072
3.7541 | -99.1932
3.7577 | -99.3222
3.7614 | -99.4536
3.7651 | -99.5575
3.7688 | -99.6319
3.7724 | -99.6626
3.7761 | -99.652
3.7798 | -99.6198
3.7835 | -99.5774
3.7871 | -99.5311
3.7908 | -99.4931
3.7945 | -99.4647
3.7982 | -99.4406
3.8018 | -99.4118
3.8055 | -99.39
3.8092 | -99.3686
3.8128 | -99.3401
3.8165 | -99.2894
3.8202 | -99.218
3.8239 | -99.1268
3.8275 | -99.0331
3.8312 | -98.9452
3.8349 | -98.8788
3.8386 | -98.8314
3.8422 | -98.8079
3.8459 | -98.8067
3.8496 | -98.8245
3.8533 | -98.8555
3.8569 | -98.8921
3.8606 | -98.922
3.8643 | -98.9505
3.8679 | -98.9599
3.8716 | -98.9601
3.8753 | -98.9616
3.879 | -98.9643
3.8826 | -98.9558
3.8863 | -98.9467
3.89 | -98.9362
3.8937 | -98.9207
3.8973 | -98.9113
3.901 | -98.9055
3.9047 | -98.8975
3.9084 | -98.896
3.912 | -98.9045
3.9157 | -98.9253
3.9194 | -98.9573
3.923 | -98.9972
3.9267 | -99.0375
3.9304 | -99.0828
3.9341 | -99.1232
3.9377 | -99.1533
3.9414 | -99.1818
3.9451 | -99.2102
3.9488 | -99.2316
3.9524 | -99.2535
3.9561 | -99.2769
3.9598 | -99.3036
3.9635 | -99.3338
3.9671 | -99.3516
3.9708 | -99.3486
3.9745 | -99.3274
3.9781 | -99.2993
3.9818 | -99.2767
3.9855 | -99.2641
3.9892 | -99.2694
3.9928 | -99.2954
3.9965 | -99.3375
4.0002 | -99.3853
4.0039 | -99.4252
4.0075 | -99.452
4.0112 | -99.4616
4.0149 | -99.4558
4.0186 | -99.4204
4.0222 | -99.3629
4.0259 | -99.2943
4.0296 | -99.2207
4.0332 | -99.1444
4.0369 | -99.0684
4.0406 | -98.9899
4.0443 | -98.9195
4.0479 | -98.8613
4.0516 | -98.8118
4.0553 | -98.7739
4.059 | -98.7508
4.0626 | -98.7496
4.0663 | -98.7817
4.07 | -98.8438
4.0736 | -98.9378
4.0773 | -99.0611
4.081 | -99.2047
4.0847 | -99.335
4.0883 | -99.4367
4.092 | -99.5059
4.0957 | -99.5285
4.0994 | -99.5168
4.103 | -99.4801
4.1067 | -99.4265
4.1104 | -99.3686
4.1141 | -99.3178
4.1177 | -99.2796
4.1214 | -99.2583
4.1251 | -99.2589
4.1287 | -99.2773
4.1324 | -99.3108
4.1361 | -99.3376
4.1398 | -99.3505
4.1434 | -99.3423
4.1471 | -99.3141
4.1508 | -99.2666
4.1545 | -99.2022
4.1581 | -99.1418
4.1618 | -99.082
4.1655 | -99.0432
4.1692 | -99.0302
4.1728 | -99.0375
4.1765 | -99.0612
4.1802 | -99.1002
4.1838 | -99.1498
4.1875 | -99.2101
4.1912 | -99.2737
4.1949 | -99.3412
4.1985 | -99.4177
4.2022 | -99.4836
4.2059 | -99.5331
4.2096 | -99.5569
4.2132 | -99.5626
4.2169 | -99.5498
4.2206 | -99.5127
4.2243 | -99.46
4.2279 | -99.3857
4.2316 | -99.3084
4.2353 | -99.2457
4.2389 | -99.2091
4.2426 | -99.1712
4.2463 | -99.1409
4.25 | -99.1225
4.2536 | -99.1181
4.2573 | -99.134
4.261 | -99.1657
4.2647 | -99.208
4.2683 | -99.2613
4.272 | -99.3235
4.2757 | -99.3823
4.2794 | -99.4137
4.283 | -99.421
4.2867 | -99.3988
4.2904 | -99.3579
4.294 | -99.3169
4.2977 | -99.2742
4.3014 | -99.2457
4.3051 | -99.2494
4.3087 | -99.2946
4.3124 | -99.3825
4.3161 | -99.5058
4.3198 | -99.6461
4.3234 | -99.7921
4.3271 | -99.9216
4.3308 | -100.0167
4.3345 | -100.0776
4.3381 | -100.11
4.3418 | -100.1069
4.3455 | -100.0821
4.3491 | -100.0523
4.3528 | -100.0212
4.3565 | -99.9968
4.3602 | -99.9741
4.3638 | -99.9572
4.3675 | -99.9547
4.3712 | -99.9669
4.3749 | -99.9877
4.3785 | -100.0215
4.3822 | -100.0642
4.3859 | -100.1103
4.3895 | -100.165
4.3932 | -100.2104
4.3969 | -100.2596
4.4006 | -100.3166
4.4042 | -100.3668
4.4079 | -100.4088
4.4116 | -100.4579
4.4153 | -100.5044
4.4189 | -100.5391
4.4226 | -100.5408
4.4263 | -100.5063
4.43 | -100.4415
4.4336 | -100.3485
4.4373 | -100.2443
4.441 | -100.1349
4.4446 | -100.0258
4.4483 | -99.9288
4.452 | -99.852
4.4557 | -99.8075
4.4593 | -99.7897
4.463 | -99.8026
4.4667 | -99.8433
4.4704 | -99.905
4.474 | -99.9871
4.4777 | -100.0712
4.4814 | -100.1395
4.4851 | -100.1917
4.4887 | -100.2175
4.4924 | -100.221
4.4961 | -100.2104
4.4997 | -100.1855
4.5034 | -100.1515
4.5071 | -100.1105
4.5108 | -100.0649
4.5144 | -100.0238
4.5181 | -99.9816
4.5218 | -99.9486
4.5255 | -99.9325
4.5291 | -99.9413
4.5328 | -99.9681
4.5365 | -100.0112
4.5402 | -100.0636
4.5438 | -100.115
4.5475 | -100.1815
4.5512 | -100.2565
4.5548 | -100.3324
4.5585 | -100.4139
4.5622 | -100.4997
4.5659 | -100.5702
4.5695 | -100.624
4.5732 | -100.6578
4.5769 | -100.6636
4.5806 | -100.6533
4.5842 | -100.6318
4.5879 | -100.6049
4.5916 | -100.5815
4.5953 | -100.5729
4.5989 | -100.5804
4.6026 | -100.6108
4.6063 | -100.6472
4.6099 | -100.6834
4.6136 | -100.705
4.6173 | -100.6927
4.621 | -100.6582
4.6246 | -100.6125
4.6283 | -100.5608
4.632 | -100.507
4.6357 | -100.4615
4.6393 | -100.4303
4.643 | -100.413
4.6467 | -100.407
4.6504 | -100.4114
4.654 | -100.4227
4.6577 | -100.4349
4.6614 | -100.4449
4.665 | -100.445
4.6687 | -100.4401
4.6724 | -100.4204
4.6761 | -100.3967
4.6797 | -100.3648
4.6834 | -100.3247
4.6871 | -100.2885
4.6908 | -100.2607
4.6944 | -100.233
4.6981 | -100.2062
4.7018 | -100.1783
4.7054 | -100.1465
4.7091 | -100.1081
4.7128 | -100.0758
4.7165 | -100.0612
4.7201 | -100.0654
4.7238 | -100.0975
4.7275 | -100.1631
4.7312 | -100.2519
4.7348 | -100.3559
4.7385 | -100.469
4.7422 | -100.576
4.7459 | -100.6516
4.7495 | -100.6875
4.7532 | -100.6728
4.7569 | -100.6087
4.7605 | -100.5142
4.7642 | -100.4169
4.7679 | -100.3443
4.7716 | -100.277
4.7752 | -100.2219
4.7789 | -100.1819
4.7826 | -100.1599
4.7863 | -100.1446
4.7899 | -100.1443
4.7936 | -100.1623
4.7973 | -100.1901
4.801 | -100.2239
4.8046 | -100.257
4.8083 | -100.2835
4.812 | -100.3066
4.8156 | -100.337
4.8193 | -100.3738
4.823 | -100.4226
4.8267 | -100.4824
4.8303 | -100.5525
4.834 | -100.6268
4.8377 | -100.701
4.8414 | -100.779
4.845 | -100.8488
4.8487 | -100.9147
4.8524 | -100.9748
4.8561 | -101.0216
4.8597 | -101.0577
4.8634 | -101.0858
4.8671 | -101.0978
4.8707 | -101.0903
4.8744 | -101.075
4.8781 | -101.0473
4.8818 | -101.02
4.8854 | -100.9851
4.8891 | -100.9302
4.8928 | -100.8676
4.8965 | -100.7996
4.9001 | -100.7327
4.9038 | -100.675
4.9075 | -100.6204
4.9112 | -100.5853
4.9148 | -100.5709
4.9185 | -100.5697
4.9222 | -100.5877
4.9258 | -100.622
4.9295 | -100.6653
4.9332 | -100.7159
4.9369 | -100.7623
4.9405 | -100.8034
4.9442 | -100.8338
4.9479 | -100.858
4.9516 | -100.8735
4.9552 | -100.8888
4.9589 | -100.8932
4.9626 | -100.8532
4.9663 | -100.8135
4.9699 | -100.7679
4.9736 | -100.7015
4.9773 | -100.5984
4.9809 | -100.4591
4.9846 | -100.3056
4.9883 | -100.1497
4.992 | -100.014
4.9956 | -99.9173
4.9993 | -99.8688
5.003 | -99.8775
5.0067 | -99.9432
5.0103 | -100.0601
5.014 | -100.2155
5.0177 | -100.39
5.0214 | -100.5703
5.025 | -100.7299
5.0287 | -100.8661
5.0324 | -100.969
5.036 | -101.042
5.0397 | -101.1093
5.0434 | -101.1188
5.0471 | -101.0935
5.0507 | -101.0644
5.0544 | -101.0382
5.0581 | -101.0214
5.0618 | -101.011
5.0654 | -101.0092
5.0691 | -101.0196
5.0728 | -101.041
5.0765 | -101.076
5.0801 | -101.1104
5.0838 | -101.1579
5.0875 | -101.212
5.0911 | -101.2657
5.0948 | -101.3242
5.0985 | -101.3726
5.1022 | -101.4169
5.1058 | -101.4477
5.1095 | -101.462
5.1132 | -101.4593
5.1169 | -101.4358
5.1205 | -101.3962
5.1242 | -101.3382
5.1279 | -101.2739
5.1315 | -101.2097
5.1352 | -101.1448
5.1389 | -101.0784
5.1426 | -101.024
5.1462 | -100.989
5.1499 | -100.9802
5.1536 | -100.9906
5.1573 | -101.0277
5.1609 | -101.0841
5.1646 | -101.1625
5.1683 | -101.2609
5.172 | -101.3662
5.1756 | -101.4752
5.1793 | -101.5774
5.183 | -101.655
5.1866 | -101.7072
5.1903 | -101.727
5.194 | -101.7068
5.1977 | -101.6457
5.2013 | -101.5633
5.205 | -101.4573
5.2087 | -101.3461
5.2124 | -101.2416
5.216 | -101.1547
5.2197 | -101.0917
5.2234 | -101.052
5.2271 | -101.0399
5.2307 | -101.0591
5.2344 | -101.1013
5.2381 | -101.1566
5.2417 | -101.221
5.2454 | -101.2911
5.2491 | -101.3502
5.2528 | -101.385
5.2564 | -101.4004
5.2601 | -101.4136
5.2638 | -101.4181
5.2675 | -101.4159
5.2711 | -101.4139
5.2748 | -101.4098
5.2785 | -101.4013
5.2822 | -101.3955
5.2858 | -101.3833
5.2895 | -101.3772
5.2932 | -101.3785
5.2968 | -101.4046
5.3005 | -101.4192
5.3042 | -101.442
5.3079 | -101.469
5.3115 | -101.4963
5.3152 | -101.5279
5.3189 | -101.5552
5.3226 | -101.5679
5.3262 | -101.5702
5.3299 | -101.5695
5.3336 | -101.5704
5.3373 | -101.5605
5.3409 | -101.5314
5.3446 | -101.4937
5.3483 | -101.4492
5.3519 | -101.4018
5.3556 | -101.3454
5.3593 | -101.2902
5.363 | -101.2444
5.3666 | -101.2097
5.3703 | -101.1795
5.374 | -101.1541
5.3777 | -101.1408
5.3813 | -101.1408
5.385 | -101.147
5.3887 | -101.1634
5.3924 | -101.1876
5.396 | -101.2175
5.3997 | -101.2475
5.4034 | -101.2777
5.407 | -101.3092
5.4107 | -101.3346
5.4144 | -101.3564
5.4181 | -101.3809
5.4217 | -101.4051
5.4254 | -101.4331
5.4291 | -101.4751
5.4328 | -101.5324
5.4364 | -101.5898
5.4401 | -101.6524
5.4438 | -101.6995
5.4474 | -101.7246
5.4511 | -101.7296
5.4548 | -101.7123
5.4585 | -101.6801
5.4621 | -101.6256
5.4658 | -101.5532
5.4695 | -101.4837
5.4732 | -101.422
5.4768 | -101.3768
5.4805 | -101.3473
5.4842 | -101.3363
5.4879 | -101.3492
5.4915 | -101.3758
5.4952 | -101.4088
5.4989 | -101.4439
5.5025 | -101.4755
5.5062 | -101.5063
5.5099 | -101.5432
5.5136 | -101.5781
5.5172 | -101.6051
5.5209 | -101.6345
5.5246 | -101.6605
5.5283 | -101.6916
5.5319 | -101.7267
5.5356 | -101.7542
5.5393 | -101.7825
5.543 | -101.8016
5.5466 | -101.8216
5.5503 | -101.8408
5.554 | -101.8515
5.5576 | -101.8451
5.5613 | -101.8277
5.565 | -101.805
5.5687 | -101.7739
5.5723 | -101.7359
5.576 | -101.7098
5.5797 | -101.6922
5.5834 | -101.6898
5.587 | -101.6989
5.5907 | -101.7276
5.5944 | -101.7572
5.5981 | -101.7877
5.6017 | -101.8292
5.6054 | -101.8719
5.6091 | -101.9022
5.6127 | -101.909
5.6164 | -101.9054
5.6201 | -101.8931
5.6238 | -101.8691
5.6274 | -101.8449
5.6311 | -101.8141
5.6348 | -101.7834
5.6385 | -101.7558
5.6421 | -101.7356
5.6458 | -101.727
5.6495 | -101.7271
5.6532 | -101.7379
5.6568 | -101.753
5.6605 | -101.7649
5.6642 | -101.7859
5.6678 | -101.8058
5.6715 | -101.8324
5.6752 | -101.8499
5.6789 | -101.8685
5.6825 | -101.8887
5.6862 | -101.9046
5.6899 | -101.916
5.6936 | -101.9269
5.6972 | -101.9368
5.7009 | -101.9422
5.7046 | -101.9465
5.7083 | -101.9429
5.7119 | -101.94
5.7156 | -101.9413
5.7193 | -101.9458
5.7229 | -101.9471
5.7266 | -101.9546
5.7303 | -101.9599
5.734 | -101.9708
5.7376 | -101.9941
5.7413 | -102.0259
5.745 | -102.0706
5.7487 | -102.1191
5.7523 | -102.167
5.756 | -102.2036
5.7597 | -102.2263
5.7633 | -102.2274
5.767 | -102.2028
5.7707 | -102.1418
5.7744 | -102.0551
5.778 | -101.9549
5.7817 | -101.8476
5.7854 | -101.7336
5.7891 | -101.6307
5.7927 | -101.5476
5.7964 | -101.4825
5.8001 | -101.4475
5.8038 | -101.4431
5.8074 | -101.4668
5.8111 | -101.5123
5.8148 | -101.5814
5.8184 | -101.67
5.8221 | -101.7682
5.8258 | -101.8513
5.8295 | -101.9326
5.8331 | -102.0014
5.8368 | -102.055
5.8405 | -102.0864
5.8442 | -102.0901
5.8478 | -102.0752
5.8515 | -102.0386
5.8552 | -101.9839
5.8589 | -101.926
5.8625 | -101.8629
5.8662 | -101.7957
5.8699 | -101.7328
5.8735 | -101.6772
5.8772 | -101.6378
5.8809 | -101.6103
5.8846 | -101.5938
5.8882 | -101.5931
5.8919 | -101.6101
5.8956 | -101.6454
5.8993 | -101.6916
5.9029 | -101.7411
5.9066 | -101.7881
5.9103 | -101.8274
5.914 | -101.8618
5.9176 | -101.8953
5.9213 | -101.9199
5.925 | -101.9353
5.9286 | -101.9446
5.9323 | -101.9533
5.936 | -101.9599
5.9397 | -101.9596
5.9433 | -101.9577
5.947 | -101.9545
5.9507 | -101.9431
5.9544 | -101.9302
5.958 | -101.9165
5.9617 | -101.8946
5.9654 | -101.8662
5.9691 | -101.8263
5.9727 | -101.787
5.9764 | -101.7534
5.9801 | -101.7204
5.9837 | -101.6958
5.9874 | -101.6798
5.9911 | -101.6741
5.9948 | -101.6801
5.9984 | -101.6977
6.0021 | -101.7291
6.0058 | -101.7682
6.0095 | -101.8092
6.0131 | -101.859
6.0168 | -101.9165
6.0205 | -101.9764
6.0242 | -102.0399
6.0278 | -102.0983
6.0315 | -102.1479
6.0352 | -102.1951
6.0388 | -102.2299
6.0425 | -102.2649
6.0462 | -102.2877
6.0499 | -102.2964
6.0535 | -102.3016
6.0572 | -102.3064
6.0609 | -102.3011
6.0646 | -102.2916
6.0682 | -102.2761
6.0719 | -102.2647
6.0756 | -102.2541
6.0793 | -102.2427
6.0829 | -102.2416
6.0866 | -102.2432
6.0903 | -102.2388
6.0939 | -102.2267
6.0976 | -102.2111
6.1013 | -102.1848
6.105 | -102.1496
6.1086 | -102.1124
6.1123 | -102.0779
6.116 | -102.0403
6.1197 | -102.0115
6.1233 | -101.9919
6.127 | -101.9747
6.1307 | -101.9772
6.1344 | -101.9845
6.138 | -102.0037
6.1417 | -102.0372
6.1454 | -102.0783
6.149 | -102.1275
6.1527 | -102.1937
6.1564 | -102.2687
6.1601 | -102.3566
6.1637 | -102.4427
6.1674 | -102.5341
6.1711 | -102.6236
6.1748 | -102.6983
6.1784 | -102.771
6.1821 | -102.8285
6.1858 | -102.8641
6.1894 | -102.8745
6.1931 | -102.8679
6.1968 | -102.8391
6.2005 | -102.783
6.2041 | -102.713
6.2078 | -102.64
6.2115 | -102.5665
6.2152 | -102.4946
6.2188 | -102.4327
6.2225 | -102.3795
6.2262 | -102.3412
6.2299 | -102.3202
6.2335 | -102.3087
6.2372 | -102.3162
6.2409 | -102.3273
6.2445 | -102.3397
6.2482 | -102.352
6.2519 | -102.3637
6.2556 | -102.3775
6.2592 | -102.3874
6.2629 | -102.398
6.2666 | -102.4118
6.2703 | -102.4247
6.2739 | -102.4558
6.2776 | -102.498
6.2813 | -102.5419
6.285 | -102.5814
6.2886 | -102.6192
6.2923 | -102.6556
6.296 | -102.6794
6.2996 | -102.6925
6.3033 | -102.6878
6.307 | -102.6676
6.3107 | -102.6412
6.3143 | -102.6116
6.318 | -102.5797
6.3217 | -102.5436
6.3254 | -102.5112
6.329 | -102.4791
6.3327 | -102.4616
6.3364 | -102.456
6.3401 | -102.4627
6.3437 | -102.4784
6.3474 | -102.5078
6.3511 | -102.5488
6.3547 | -102.5974
6.3584 | -102.6613
6.3621 | -102.7165
6.3658 | -102.7716
6.3694 | -102.8249
6.3731 | -102.8787
6.3768 | -102.9332
6.3805 | -102.9805
6.3841 | -103.0079
6.3878 | -103.0187
6.3915 | -103.0213
6.3952 | -103.009
6.3988 | -102.9885
6.4025 | -102.9657
6.4062 | -102.9308
6.4098 | -102.8887
6.4135 | -102.8474
6.4172 | -102.807
6.4209 | -102.771
6.4245 | -102.7301
6.4282 | -102.6872
6.4319 | -102.6501
6.4356 | -102.6138
6.4392 | -102.5876
6.4429 | -102.5721
6.4466 | -102.5611
6.4502 | -102.5535
6.4539 | -102.5493
6.4576 | -102.5489
6.4613 | -102.5483
6.4649 | -102.5474
6.4686 | -102.5479
6.4723 | -102.5449
6.476 | -102.5438
6.4796 | -102.5496
6.4833 | -102.5562
6.487 | -102.5625
6.4907 | -102.5616
6.4943 | -102.5559
6.498 | -102.5469
6.5017 | -102.5308
6.5053 | -102.5108
6.509 | -102.489
6.5127 | -102.4687
6.5164 | -102.4562
6.52 | -102.4511
6.5237 | -102.4625
6.5274 | -102.4879
6.5311 | -102.5195
6.5347 | -102.5637
6.5384 | -102.6142
6.5421 | -102.6636
6.5458 | -102.721
6.5494 | -102.784
6.5531 | -102.8439
6.5568 | -102.8919
6.5604 | -102.9369
6.5641 | -102.9771
6.5678 | -103.0031
6.5715 | -103.0199
6.5751 | -103.0196
6.5788 | -103.0038
6.5825 | -102.9869
6.5862 | -102.9658
6.5898 | -102.9418
6.5935 | -102.9156
6.5972 | -102.8825
6.6009 | -102.8457
6.6045 | -102.8099
6.6082 | -102.7759
6.6119 | -102.74
6.6155 | -102.7024
6.6192 | -102.6642
6.6229 | -102.6283
6.6266 | -102.5998
6.6302 | -102.5762
6.6339 | -102.5567
6.6376 | -102.5401
6.6413 | -102.5385
6.6449 | -102.5404
6.6486 | -102.544
6.6523 | -102.5502
6.656 | -102.562
6.6596 | -102.5778
6.6633 | -102.592
6.667 | -102.6085
6.6706 | -102.6295
6.6743 | -102.6514
6.678 | -102.6695
6.6817 | -102.686
6.6853 | -102.7054
6.689 | -102.7299
6.6927 | -102.7589
6.6964 | -102.7978
6.7 | -102.8473
6.7037 | -102.9032
6.7074 | -102.9662
6.7111 | -103.0294
6.7147 | -103.0937
6.7184 | -103.1512
6.7221 | -103.2028
6.7257 | -103.2435
6.7294 | -103.2685
6.7331 | -103.2891
6.7368 | -103.2976
6.7404 | -103.2864
6.7441 | -103.2709
6.7478 | -103.2494
6.7515 | -103.2286
6.7551 | -103.2028
6.7588 | -103.1663
6.7625 | -103.1255
6.7662 | -103.0812
6.7698 | -103.0343
6.7735 | -102.9925
6.7772 | -102.9529
6.7808 | -102.9111
6.7845 | -102.8712
6.7882 | -102.8392
6.7919 | -102.8121
6.7955 | -102.7902
6.7992 | -102.777
6.8029 | -102.7767
6.8066 | -102.7851
6.8102 | -102.81
6.8139 | -102.8364
6.8176 | -102.8711
6.8213 | -102.9103
6.8249 | -102.9487
6.8286 | -102.9914
6.8323 | -103.0253
6.8359 | -103.0519
6.8396 | -103.073
6.8433 | -103.0833
6.847 | -103.086
6.8506 | -103.0793
6.8543 | -103.0712
6.858 | -103.0649
6.8617 | -103.0557
6.8653 | -103.049
6.869 | -103.0495
6.8727 | -103.0544
6.8763 | -103.065
6.88 | -103.0889
6.8837 | -103.1178
6.8874 | -103.1564
6.891 | -103.1887
6.8947 | -103.2235
6.8984 | -103.2526
6.9021 | -103.2746
6.9057 | -103.2894
6.9094 | -103.2979
6.9131 | -103.302
6.9168 | -103.3048
6.9204 | -103.2988
6.9241 | -103.289
6.9278 | -103.2798
6.9314 | -103.2686
6.9351 | -103.2583
6.9388 | -103.2493
6.9425 | -103.2334
6.9461 | -103.1975
6.9498 | -103.1385
6.9535 | -103.096
6.9572 | -103.0459
6.9608 | -102.9814
6.9645 | -102.9047
6.9682 | -102.8164
6.9719 | -102.7172
6.9755 | -102.6164
6.9792 | -102.5216
6.9829 | -102.4262
6.9865 | -102.3413
6.9902 | -102.2699
6.9939 | -102.2157
6.9976 | -102.1852
7.0012 | -102.1791
7.0049 | -102.1994
7.0086 | -102.242
7.0123 | -102.3014
7.0159 | -102.3733
7.0196 | -102.4548
7.0233 | -102.5385
7.027 | -102.6186
7.0306 | -102.6887
7.0343 | -102.7492
7.038 | -102.799
7.0416 | -102.8335
7.0453 | -102.8603
7.049 | -102.8849
7.0527 | -102.911
7.0563 | -102.9347
7.06 | -102.9222
7.0637 | -102.9117
7.0674 | -102.9066
7.071 | -102.9105
7.0747 | -102.916
7.0784 | -102.9263
7.0821 | -102.9364
7.0857 | -102.9491
7.0894 | -102.969
7.0931 | -102.9902
7.0967 | -103.0147
7.1004 | -103.0423
7.1041 | -103.0727
7.1078 | -103.1066
7.1114 | -103.1376
7.1151 | -103.1681
7.1188 | -103.1981
7.1225 | -103.2238
7.1261 | -103.2424
7.1298 | -103.251
7.1335 | -103.2606
7.1372 | -103.2637
7.1408 | -103.2642
7.1445 | -103.2586
7.1482 | -103.2511
7.1518 | -103.2384
7.1555 | -103.2234
7.1592 | -103.2118
7.1629 | -103.2056
7.1665 | -103.2027
7.1702 | -103.2002
7.1739 | -103.2044
7.1776 | -103.212
7.1812 | -103.2255
7.1849 | -103.2476
7.1886 | -103.2804
7.1922 | -103.314
7.1959 | -103.3504
7.1996 | -103.3909
7.2033 | -103.4289
7.2069 | -103.4595
7.2106 | -103.4814
7.2143 | -103.494
7.218 | -103.507
7.2216 | -103.5172
7.2253 | -103.5281
7.229 | -103.5339
7.2327 | -103.5361
7.2363 | -103.5396
7.24 | -103.5314
7.2437 | -103.5219
7.2473 | -103.5104
7.251 | -103.5051
7.2547 | -103.5
7.2584 | -103.4911
7.262 | -103.4776
7.2657 | -103.4584
7.2694 | -103.4386
7.2731 | -103.419
7.2767 | -103.4001
7.2804 | -103.3798
7.2841 | -103.3632
7.2878 | -103.352
7.2914 | -103.3486
7.2951 | -103.3509
7.2988 | -103.3631
7.3024 | -103.383
7.3061 | -103.406
7.3098 | -103.4332
7.3135 | -103.4658
7.3171 | -103.4981
7.3208 | -103.5301
7.3245 | -103.551
7.3282 | -103.5683
7.3318 | -103.5822
7.3355 | -103.5957
7.3392 | -103.5971
7.3429 | -103.5895
7.3465 | -103.5715
7.3502 | -103.5423
7.3539 | -103.5063
7.3575 | -103.4596
7.3612 | -103.4126
7.3649 | -103.3699
7.3686 | -103.3307
7.3722 | -103.3014
7.3759 | -103.2715
7.3796 | -103.2559
7.3833 | -103.255
7.3869 | -103.2631
7.3906 | -103.2772
7.3943 | -103.3032
7.398 | -103.3365
7.4016 | -103.3783
7.4053 | -103.4208
7.409 | -103.4692
7.4126 | -103.5251
7.4163 | -103.5707
7.42 | -103.6196
7.4237 | -103.6589
7.4273 | -103.6913
7.431 | -103.7181
7.4347 | -103.7414
7.4384 | -103.7522
7.442 | -103.7505
7.4457 | -103.738
7.4494 | -103.7135
7.4531 | -103.6814
7.4567 | -103.6414
7.4604 | -103.5943
7.4641 | -103.5408
7.4677 | -103.4837
7.4714 | -103.43
7.4751 | -103.3807
7.4788 | -103.3354
7.4824 | -103.2997
7.4861 | -103.2764
7.4898 | -103.2574
7.4935 | -103.2372
7.4971 | -103.224
7.5008 | -103.2158
7.5045 | -103.2073
7.5081 | -103.1955
7.5118 | -103.1871
7.5155 | -103.1808
7.5192 | -103.1694
7.5228 | -103.1578
7.5265 | -103.145
7.5302 | -103.1251
7.5339 | -103.1032
7.5375 | -103.0871
7.5412 | -103.0725
7.5449 | -103.0615
7.5486 | -103.0529
7.5522 | -103.0521
7.5559 | -103.0584
7.5596 | -103.0747
7.5632 | -103.102
7.5669 | -103.1339
7.5706 | -103.1676
7.5743 | -103.2112
7.5779 | -103.2638
7.5816 | -103.3204
7.5853 | -103.3755
7.589 | -103.4278
7.5926 | -103.4795
7.5963 | -103.5231
7.6 | -103.5625
7.6037 | -103.5962
7.6073 | -103.6259
7.611 | -103.6444
7.6147 | -103.6531
7.6183 | -103.6514
7.622 | -103.6452
7.6257 | -103.6281
7.6294 | -103.61
7.633 | -103.5927
7.6367 | -103.5691
7.6404 | -103.5573
7.6441 | -103.558
7.6477 | -103.5597
7.6514 | -103.5664
7.6551 | -103.5783
7.6588 | -103.5939
7.6624 | -103.612
7.6661 | -103.6391
7.6698 | -103.6719
7.6734 | -103.702
7.6771 | -103.7328
7.6808 | -103.7629
7.6845 | -103.7954
7.6881 | -103.8128
7.6918 | -103.8215
7.6955 | -103.8304
7.6992 | -103.836
7.7028 | -103.8354
7.7065 | -103.8252
7.7102 | -103.8062
7.7139 | -103.7782
7.7175 | -103.7423
7.7212 | -103.7024
7.7249 | -103.6573
7.7285 | -103.6166
7.7322 | -103.5804
7.7359 | -103.5426
7.7396 | -103.5068
7.7432 | -103.472
7.7469 | -103.4411
7.7506 | -103.4185
7.7543 | -103.4003
7.7579 | -103.3895
7.7616 | -103.3946
7.7653 | -103.4041
7.769 | -103.4281
7.7726 | -103.4611
7.7763 | -103.5058
7.78 | -103.5552
7.7836 | -103.6136
7.7873 | -103.6805
7.791 | -103.7463
7.7947 | -103.8028
7.7983 | -103.854
7.802 | -103.9074
7.8057 | -103.9487
7.8094 | -103.9754
7.813 | -103.9884
7.8167 | -103.9982
7.8204 | -103.9993
7.8241 | -103.9916
7.8277 | -103.9793
7.8314 | -103.9621
7.8351 | -103.9408
7.8387 | -103.9205
7.8424 | -103.8954
7.8461 | -103.8728
7.8498 | -103.8596
7.8534 | -103.8532
7.8571 | -103.8509
7.8608 | -103.8547
7.8645 | -103.8608
7.8681 | -103.872
7.8718 | -103.8789
7.8755 | -103.8818
7.8792 | -103.8843
7.8828 | -103.883
7.8865 | -103.8819
7.8902 | -103.87
7.8938 | -103.8514
7.8975 | -103.8329
7.9012 | -103.814
7.9049 | -103.8035
7.9085 | -103.7963
7.9122 | -103.7802
7.9159 | -103.7627
7.9196 | -103.7513
7.9232 | -103.743
7.9269 | -103.738
7.9306 | -103.7259
7.9342 | -103.7177
7.9379 | -103.7166
7.9416 | -103.7188
7.9453 | -103.7239
7.9489 | -103.7434
7.9526 | -103.7581
7.9563 | -103.7747
7.96 | -103.7933
7.9636 | -103.8112
7.9673 | -103.8334
7.971 | -103.857
7.9747 | -103.8822
7.9783 | -103.9072
7.982 | -103.931
7.9857 | -103.9493
7.9893 | -103.9731
7.993 | -103.9963
7.9967 | -104.025
8.0004 | -104.0527
8.004 | -104.0816
8.0077 | -104.1196
8.0114 | -104.1578
8.0151 | -104.1877
8.0187 | -104.2182
8.0224 | -104.2423
8.0261 | -104.2676
8.0298 | -104.2936
8.0334 | -104.3146
8.0371 | -104.3313
8.0408 | -104.3423
8.0444 | -104.3515
8.0481 | -104.3569
8.0518 | -104.3557
8.0555 | -104.3471
8.0591 | -104.3291
8.0628 | -104.3029
8.0665 | -104.2684
8.0702 | -104.2274
8.0738 | -104.1825
8.0775 | -104.1332
8.0812 | -104.0785
8.0849 | -104.0223
8.0885 | -103.9657
8.0922 | -103.9152
8.0959 | -103.8652
8.0995 | -103.8179
8.1032 | -103.773
8.1069 | -103.7272
8.1106 | -103.6886
8.1142 | -103.6573
8.1179 | -103.6268
8.1216 | -103.6003
8.1253 | -103.5792
8.1289 | -103.5621
8.1326 | -103.5523
8.1363 | -103.5515
8.14 | -103.553
8.1436 | -103.5581
8.1473 | -103.5655
8.151 | -103.5765
8.1546 | -103.5852
8.1583 | -103.5956
8.162 | -103.61
8.1657 | -103.6253
8.1693 | -103.6443
8.173 | -103.6587
8.1767 | -103.6733
8.1804 | -103.692
8.184 | -103.7126
8.1877 | -103.732
8.1914 | -103.7494
8.1951 | -103.7642
8.1987 | -103.7823
8.2024 | -103.7976
8.2061 | -103.8107
8.2097 | -103.8169
8.2134 | -103.8204
8.2171 | -103.8246
8.2208 | -103.826
8.2244 | -103.8303
8.2281 | -103.8392
8.2318 | -103.8444
8.2355 | -103.8506
8.2391 | -103.8634
8.2428 | -103.8852
8.2465 | -103.9107
8.2501 | -103.9417
8.2538 | -103.983
8.2575 | -104.0342
8.2612 | -104.096
8.2648 | -104.1629
8.2685 | -104.2292
8.2722 | -104.2941
8.2759 | -104.3648
8.2795 | -104.4348
8.2832 | -104.4928
8.2869 | -104.5427
8.2906 | -104.5913
8.2942 | -104.6306
8.2979 | -104.6578
8.3016 | -104.6723
8.3052 | -104.6749
8.3089 | -104.6617
8.3126 | -104.6359
8.3163 | -104.6072
8.3199 | -104.5627
8.3236 | -104.5096
8.3273 | -104.4543
8.331 | -104.3991
8.3346 | -104.3437
8.3383 | -104.2897
8.342 | -104.2405
8.3457 | -104.1963
8.3493 | -104.1548
8.353 | -104.1164
8.3567 | -104.092
8.3603 | -104.074
8.364 | -104.0582
8.3677 | -104.0526
8.3714 | -104.0549
8.375 | -104.0575
8.3787 | -104.0639
8.3824 | -104.0798
8.3861 | -104.1052
8.3897 | -104.1387
8.3934 | -104.1738
8.3971 | -104.2074
8.4008 | -104.2415
8.4044 | -104.2813
8.4081 | -104.3203
8.4118 | -104.355
8.4154 | -104.387
8.4191 | -104.4104
8.4228 | -104.4288
8.4265 | -104.4443
8.4301 | -104.4568
8.4338 | -104.4595
8.4375 | -104.4505
8.4412 | -104.441
8.4448 | -104.4323
8.4485 | -104.416
8.4522 | -104.3927
8.4559 | -104.3734
8.4595 | -104.3608
8.4632 | -104.3507
8.4669 | -104.3397
8.4705 | -104.3356
8.4742 | -104.3367
8.4779 | -104.3409
8.4816 | -104.3367
8.4852 | -104.3386
8.4889 | -104.3429
8.4926 | -104.3439
8.4963 | -104.3447
8.4999 | -104.3506
8.5036 | -104.3568
8.5073 | -104.3587
8.511 | -104.3623
8.5146 | -104.3699
8.5183 | -104.3811
8.522 | -104.3925
8.5256 | -104.4013
8.5293 | -104.4113
8.533 | -104.4197
8.5367 | -104.4275
8.5403 | -104.4377
8.544 | -104.4466
8.5477 | -104.4525
8.5514 | -104.455
8.555 | -104.4564
8.5587 | -104.4557
8.5624 | -104.4533
8.5661 | -104.4483
8.5697 | -104.4425
8.5734 | -104.437
8.5771 | -104.4356
8.5807 | -104.4309
8.5844 | -104.4293
8.5881 | -104.4306
8.5918 | -104.4294
8.5954 | -104.4217
8.5991 | -104.4168
8.6028 | -104.4189
8.6065 | -104.4211
8.6101 | -104.419
8.6138 | -104.4156
8.6175 | -104.4152
8.6211 | -104.416
8.6248 | -104.4123
8.6285 | -104.406
8.6322 | -104.3994
8.6358 | -104.3922
8.6395 | -104.3839
8.6432 | -104.3719
8.6469 | -104.3572
8.6505 | -104.3422
8.6542 | -104.329
8.6579 | -104.315
8.6616 | -104.3002
8.6652 | -104.2838
8.6689 | -104.2646
8.6726 | -104.2458
8.6762 | -104.2307
8.6799 | -104.216
8.6836 | -104.2006
8.6873 | -104.1902
8.6909 | -104.1805
8.6946 | -104.1688
8.6983 | -104.156
8.702 | -104.1504
8.7056 | -104.1476
8.7093 | -104.1441
8.713 | -104.1434
8.7167 | -104.1483
8.7203 | -104.1587
8.724 | -104.1676
8.7277 | -104.1751
8.7313 | -104.1748
8.735 | -104.1761
8.7387 | -104.1855
8.7424 | -104.1901
8.746 | -104.1921
8.7497 | -104.1941
8.7534 | -104.1985
8.7571 | -104.2004
8.7607 | -104.2028
8.7644 | -104.2049
8.7681 | -104.2078
8.7718 | -104.2129
8.7754 | -104.2196
8.7791 | -104.226
8.7828 | -104.2323
8.7864 | -104.2389
8.7901 | -104.2448
8.7938 | -104.2478
8.7975 | -104.2522
8.8011 | -104.2586
8.8048 | -104.2641
8.8085 | -104.273
8.8122 | -104.2801
8.8158 | -104.2855
8.8195 | -104.2908
8.8232 | -104.2975
8.8269 | -104.3022
8.8305 | -104.3013
8.8342 | -104.2975
8.8379 | -104.2968
8.8415 | -104.2943
8.8452 | -104.2898
8.8489 | -104.286
8.8526 | -104.2855
8.8562 | -104.2814
8.8599 | -104.2784
8.8636 | -104.2745
8.8673 | -104.2645
8.8709 | -104.2606
8.8746 | -104.2584
8.8783 | -104.251
8.882 | -104.2422
8.8856 | -104.2343
8.8893 | -104.2227
8.893 | -104.2158
8.8966 | -104.2056
8.9003 | -104.1928
8.904 | -104.1798
8.9077 | -104.1682
8.9113 | -104.1577
8.915 | -104.1423
8.9187 | -104.1255
8.9224 | -104.1086
8.926 | -104.0863
8.9297 | -104.0582
8.9334 | -104.0282
8.937 | -103.9792
8.9407 | -103.9346
8.9444 | -103.8942
8.9481 | -103.8522
8.9517 | -103.8117
8.9554 | -103.7711
8.9591 | -103.7333
8.9628 | -103.698
8.9664 | -103.6669
8.9701 | -103.6412
8.9738 | -103.6232
8.9775 | -103.6117
8.9811 | -103.6109
8.9848 | -103.6168
8.9885 | -103.63
8.9921 | -103.6515
8.9958 | -103.681
8.9995 | -103.7112
9.0032 | -103.7433
9.0068 | -103.7939
9.0105 | -103.838
9.0142 | -103.8835
9.0179 | -103.9313
9.0215 | -103.9802
9.0252 | -104.0272
9.0289 | -104.0728
9.0326 | -104.1164
9.0362 | -104.1597
9.0399 | -104.1997
9.0436 | -104.236
9.0472 | -104.2729
9.0509 | -104.3056
9.0546 | -104.3298
9.0583 | -104.35
9.0619 | -104.3702
9.0656 | -104.3842
9.0693 | -104.4051
9.073 | -104.4142
9.0766 | -104.4128
9.0803 | -104.4087
9.084 | -104.4019
9.0877 | -104.3952
9.0913 | -104.3879
9.095 | -104.3834
9.0987 | -104.3779
9.1023 | -104.3712
9.106 | -104.3615
9.1097 | -104.3521
9.1134 | -104.3386
9.117 | -104.3265
9.1207 | -104.3163
9.1244 | -104.3044
9.1281 | -104.2874
9.1317 | -104.2741
9.1354 | -104.2668
9.1391 | -104.2521
9.1428 | -104.236
9.1464 | -104.2252
9.1501 | -104.2142
9.1538 | -104.2036
9.1574 | -104.1947
9.1611 | -104.1849
9.1648 | -104.183
9.1685 | -104.1811
9.1721 | -104.1804
9.1758 | -104.1862
9.1795 | -104.1942
9.1832 | -104.2042
9.1868 | -104.2153
9.1905 | -104.2307
9.1942 | -104.2487
9.1979 | -104.2648
9.2015 | -104.282
9.2052 | -104.3022
9.2089 | -104.3212
9.2125 | -104.3412
9.2162 | -104.3581
9.2199 | -104.373
9.2236 | -104.3821
9.2272 | -104.3872
9.2309 | -104.391
9.2346 | -104.397
9.2383 | -104.4052
9.2419 | -104.4177
9.2456 | -104.4286
9.2493 | -104.4397
9.253 | -104.4529
9.2566 | -104.4648
9.2603 | -104.4734
9.264 | -104.4876
9.2676 | -104.4998
9.2713 | -104.5063
9.275 | -104.5134
9.2787 | -104.5166
9.2823 | -104.515
9.286 | -104.5124
9.2897 | -104.5082
9.2934 | -104.5021
9.297 | -104.4951
9.3007 | -104.4821
9.3044 | -104.4699
9.308 | -104.4561
9.3117 | -104.4402
9.3154 | -104.4303
9.3191 | -104.4251
9.3227 | -104.4237
9.3264 | -104.422
9.3301 | -104.4184
9.3338 | -104.4175
9.3374 | -104.4219
9.3411 | -104.4282
9.3448 | -104.4383
9.3485 | -104.4516
9.3521 | -104.4715
9.3558 | -104.4893
9.3595 | -104.5123
9.3631 | -104.5404
9.3668 | -104.5622
9.3705 | -104.5801
9.3742 | -104.5958
9.3778 | -104.6105
9.3815 | -104.625
9.3852 | -104.6378
9.3889 | -104.6489
9.3925 | -104.661
9.3962 | -104.677
9.3999 | -104.6864
9.4036 | -104.6912
9.4072 | -104.6953
9.4109 | -104.6978
9.4146 | -104.6994
9.4182 | -104.7003
9.4219 | -104.7001
9.4256 | -104.6999
9.4293 | -104.7049
9.4329 | -104.7125
9.4366 | -104.7221
9.4403 | -104.7307
9.444 | -104.7338
9.4476 | -104.7357
9.4513 | -104.7351
9.455 | -104.7328
9.4587 | -104.7284
9.4623 | -104.7243
9.466 | -104.7178
9.4697 | -104.7029
9.4733 | -104.6843
9.477 | -104.6661
9.4807 | -104.6448
9.4844 | -104.6193
9.488 | -104.5925
9.4917 | -104.5676
9.4954 | -104.5417
9.4991 | -104.5152
9.5027 | -104.4973
9.5064 | -104.4807
9.5101 | -104.4684
9.5138 | -104.4608
9.5174 | -104.4572
9.5211 | -104.4588
9.5248 | -104.4647
9.5284 | -104.471
9.5321 | -104.4795
9.5358 | -104.5033
9.5395 | -104.521
9.5431 | -104.5414
9.5468 | -104.5579
9.5505 | -104.576
9.5542 | -104.5893
9.5578 | -104.6022
9.5615 | -104.6129
9.5652 | -104.6188
9.5689 | -104.6176
9.5725 | -104.6088
9.5762 | -104.5992
9.5799 | -104.5896
9.5835 | -104.58
9.5872 | -104.5651
9.5909 | -104.5469
9.5946 | -104.5299
9.5982 | -104.514
9.6019 | -104.4974
9.6056 | -104.4856
9.6093 | -104.4822
9.6129 | -104.4818
9.6166 | -104.485
9.6203 | -104.4928
9.624 | -104.5096
9.6276 | -104.5315
9.6313 | -104.5564
9.635 | -104.5856
9.6386 | -104.6214
9.6423 | -104.6611
9.646 | -104.7032
9.6497 | -104.7448
9.6533 | -104.7868
9.657 | -104.8329
9.6607 | -104.8787
9.6644 | -104.9203
9.668 | -104.9644
9.6717 | -105.0044
9.6754 | -105.0341
9.679 | -105.0574
9.6827 | -105.0721
9.6864 | -105.081
9.6901 | -105.0886
9.6937 | -105.0859
9.6974 | -105.0793
9.7011 | -105.0671
9.7048 | -105.0507
9.7084 | -105.0292
9.7121 | -105.0077
9.7158 | -104.986
9.7195 | -104.9613
9.7231 | -104.9334
9.7268 | -104.9037
9.7305 | -104.8789
9.7341 | -104.8546
9.7378 | -104.8295
9.7415 | -104.8043
9.7452 | -104.7831
9.7488 | -104.7632
9.7525 | -104.7403
9.7562 | -104.7216
9.7599 | -104.7059
9.7635 | -104.6943
9.7672 | -104.6802
9.7709 | -104.6704
9.7746 | -104.6657
9.7782 | -104.6624
9.7819 | -104.6598
9.7856 | -104.6596
9.7892 | -104.6664
9.7929 | -104.6724
9.7966 | -104.6786
9.8003 | -104.6879
9.8039 | -104.6966
9.8076 | -104.7085
9.8113 | -104.7209
9.815 | -104.735
9.8186 | -104.7513
9.8223 | -104.7614
9.826 | -104.7714
9.8297 | -104.7786
9.8333 | -104.7839
9.837 | -104.7892
9.8407 | -104.7889
9.8443 | -104.7867
9.848 | -104.784
9.8517 | -104.7864
9.8554 | -104.7895
9.859 | -104.7946
9.8627 | -104.7999
9.8664 | -104.8102
9.8701 | -104.8272
9.8737 | -104.8429
9.8774 | -104.8614
9.8811 | -104.8825
9.8848 | -104.9086
9.8884 | -104.934
9.8921 | -104.9568
9.8958 | -104.9833
9.8994 | -105.0113
9.9031 | -105.0369
9.9068 | -105.0611
9.9105 | -105.0893
9.9141 | -105.1131
9.9178 | -105.1344
9.9215 | -105.1516
9.9252 | -105.1647
9.9288 | -105.1715
9.9325 | -105.1737
9.9362 | -105.1799
9.9398 | -105.1805
9.9435 | -105.1723
9.9472 | -105.1673
9.9509 | -105.1607
9.9545 | -105.1492
9.9582 | -105.1374
9.9619 | -105.1248
9.9656 | -105.1112
9.9692 | -105.1006
9.9729 | -105.0893
9.9766 | -105.075
9.9803 | -105.0682
9.9839 | -105.0668
9.9876 | -105.0678
9.9913 | -105.0702
9.9949 | -105.0735
9.9986 | -105.0807
10.0023 | -105.0947
10.006 | -105.1103
10.0096 | -105.1253
10.0133 | -105.139
10.017 | -105.1523
10.0207 | -105.1671
10.0243 | -105.1796
10.028 | -105.194
10.0317 | -105.2024
10.0354 | -105.2076
10.039 | -105.2096
10.0427 | -105.2008
10.0464 | -105.1886
10.05 | -105.1745
10.0537 | -105.1528
10.0574 | -105.1275
10.0611 | -105.0968
10.0647 | -105.0643
10.0684 | -105.0287
10.0721 | -105.0003
10.0758 | -104.9752
10.0794 | -104.9512
10.0831 | -104.9249
10.0868 | -104.9025
10.0905 | -104.8871
10.0941 | -104.8746
10.0978 | -104.8696
10.1015 | -104.8711
10.1051 | -104.8752
10.1088 | -104.8841
10.1125 | -104.8982
10.1162 | -104.9117
10.1198 | -104.9281
10.1235 | -104.9479
10.1272 | -104.9613
10.1309 | -104.9742
10.1345 | -104.9868
10.1382 | -104.9953
10.1419 | -105.0042
10.1456 | -105.0104
10.1492 | -105.0126
10.1529 | -105.014
10.1566 | -105.0139
10.1602 | -105.0087
10.1639 | -105.0042
10.1676 | -104.9958
10.1713 | -104.9852
10.1749 | -104.9713
10.1786 | -104.9608
10.1823 | -104.9513
10.186 | -104.9399
10.1896 | -104.9286
10.1933 | -104.9153
10.197 | -104.9004
10.2007 | -104.8877
10.2043 | -104.8756
10.208 | -104.8648
10.2117 | -104.857
10.2153 | -104.8519
10.219 | -104.8468
10.2227 | -104.8429
10.2264 | -104.8437
10.23 | -104.8476
10.2337 | -104.8572
10.2374 | -104.8732
10.2411 | -104.8919
10.2447 | -104.9124
10.2484 | -104.9359
10.2521 | -104.9635
10.2558 | -104.9985
10.2594 | -105.0387
10.2631 | -105.0801
10.2668 | -105.1222
10.2704 | -105.1642
10.2741 | -105.2
10.2778 | -105.2403
10.2815 | -105.2795
10.2851 | -105.3124
10.2888 | -105.342
10.2925 | -105.3638
10.2962 | -105.3829
10.2998 | -105.3926
10.3035 | -105.4004
10.3072 | -105.4054
10.3109 | -105.4108
10.3145 | -105.4086
10.3182 | -105.4027
10.3219 | -105.3941
10.3255 | -105.3843
10.3292 | -105.3725
10.3329 | -105.3594
10.3366 | -105.3459
10.3402 | -105.3313
10.3439 | -105.3212
10.3476 | -105.3104
10.3513 | -105.3011
10.3549 | -105.2958
10.3586 | -105.2923
10.3623 | -105.2917
10.3659 | -105.2933
10.3696 | -105.3033
10.3733 | -105.3191
10.377 | -105.3384
10.3806 | -105.3592
10.3843 | -105.3822
10.388 | -105.4015
10.3917 | -105.4228
10.3953 | -105.4465
10.399 | -105.471
10.4027 | -105.4949
10.4064 | -105.5102
10.41 | -105.5189
10.4137 | -105.5308
10.4174 | -105.5416
10.421 | -105.5455
10.4247 | -105.5525
10.4284 | -105.5519
10.4321 | -105.5483
10.4357 | -105.5426
10.4394 | -105.5347
10.4431 | -105.5233
10.4468 | -105.5063
10.4504 | -105.4882
10.4541 | -105.4659
10.4578 | -105.4396
10.4615 | -105.4097
10.4651 | -105.3811
10.4688 | -105.3541
10.4725 | -105.3238
10.4761 | -105.2914
10.4798 | -105.2635
10.4835 | -105.2351
10.4872 | -105.2065
10.4908 | -105.1854
10.4945 | -105.1675
10.4982 | -105.1517
10.5019 | -105.1391
10.5055 | -105.1285
10.5092 | -105.1185
10.5129 | -105.1086
10.5166 | -105.1042
10.5202 | -105.1017
10.5239 | -105.1009
10.5276 | -105.1013
10.5312 | -105.1096
10.5349 | -105.1194
10.5386 | -105.1344
10.5423 | -105.1529
10.5459 | -105.1693
10.5496 | -105.1885
10.5533 | -105.2067
10.557 | -105.2279
10.5606 | -105.2457
10.5643 | -105.2594
10.568 | -105.2734
10.5717 | -105.2898
10.5753 | -105.3019
10.579 | -105.3134
10.5827 | -105.3228
10.5863 | -105.33
10.59 | -105.3386
10.5937 | -105.3434
10.5974 | -105.3464
10.601 | -105.3524
10.6047 | -105.3555
10.6084 | -105.3591
10.6121 | -105.3659
10.6157 | -105.3724
10.6194 | -105.3754
10.6231 | -105.3757
10.6268 | -105.3821
10.6304 | -105.3914
10.6341 | -105.3973
10.6378 | -105.4028
10.6414 | -105.4117
10.6451 | -105.4172
10.6488 | -105.4194
10.6525 | -105.425
10.6561 | -105.4289
10.6598 | -105.4293
10.6635 | -105.4337
10.6672 | -105.4367
10.6708 | -105.4329
10.6745 | -105.4281
10.6782 | -105.4241
10.6818 | -105.4166
10.6855 | -105.4052
10.6892 | -105.3928
10.6929 | -105.3801
10.6965 | -105.3638
10.7002 | -105.352
10.7039 | -105.339
10.7076 | -105.3286
10.7112 | -105.3172
10.7149 | -105.3109
10.7186 | -105.3057
10.7223 | -105.2964
10.7259 | -105.2908
10.7296 | -105.2866
10.7333 | -105.2811
10.7369 | -105.2753
10.7406 | -105.2715
10.7443 | -105.2684
10.748 | -105.2658
10.7516 | -105.2621
10.7553 | -105.2571
10.759 | -105.2546
10.7627 | -105.2529
10.7663 | -105.2477
10.77 | -105.2401
10.7737 | -105.2349
10.7774 | -105.2365
10.781 | -105.2392
10.7847 | -105.2379
10.7884 | -105.2334
10.792 | -105.2303
10.7957 | -105.2319
10.7994 | -105.232
10.8031 | -105.2327
10.8067 | -105.2339
10.8104 | -105.2321
10.8141 | -105.2301
10.8178 | -105.2318
10.8214 | -105.2299
10.8251 | -105.2267
10.8288 | -105.2307
10.8325 | -105.2316
10.8361 | -105.2318
10.8398 | -105.2357
10.8435 | -105.2399
10.8471 | -105.2464
10.8508 | -105.2534
10.8545 | -105.2664
10.8582 | -105.2795
10.8618 | -105.2907
10.8655 | -105.3044
10.8692 | -105.3233
10.8729 | -105.3439
10.8765 | -105.364
10.8802 | -105.3879
10.8839 | -105.4089
10.8876 | -105.431
10.8912 | -105.4533
10.8949 | -105.4799
10.8986 | -105.5041
10.9022 | -105.5224
10.9059 | -105.5428
10.9096 | -105.5628
10.9133 | -105.5824
10.9169 | -105.5965
10.9206 | -105.6023
10.9243 | -105.6001
10.928 | -105.5988
10.9316 | -105.5948
10.9353 | -105.584
10.939 | -105.5704
10.9427 | -105.5534
10.9463 | -105.5358
10.95 | -105.5138
10.9537 | -105.486
10.9573 | -105.4549
10.961 | -105.4156
10.9647 | -105.3752
10.9684 | -105.3393
10.972 | -105.3039
10.9757 | -105.2636
10.9794 | -105.2298
10.9831 | -105.1994
10.9867 | -105.1672
10.9904 | -105.1369
10.9941 | -105.1115
10.9977 | -105.0898
11.0014 | -105.0716
11.0051 | -105.0586
11.0088 | -105.0473
11.0124 | -105.0354
11.0161 | -105.032
11.0198 | -105.0362
11.0235 | -105.0425
11.0271 | -105.0497
11.0308 | -105.0591
11.0345 | -105.0682
11.0382 | -105.0841
11.0418 | -105.1067
11.0455 | -105.135
11.0492 | -105.162
11.0528 | -105.1892
11.0565 | -105.2201
11.0602 | -105.254
11.0639 | -105.2851
11.0675 | -105.3137
11.0712 | -105.3385
11.0749 | -105.3655
11.0786 | -105.3928
11.0822 | -105.4177
11.0859 | -105.4317
11.0896 | -105.4396
11.0933 | -105.4441
11.0969 | -105.4458
11.1006 | -105.4447
11.1043 | -105.4405
11.1079 | -105.4331
11.1116 | -105.4263
11.1153 | -105.4208
11.119 | -105.4174
11.1226 | -105.4119
11.1263 | -105.4005
11.13 | -105.3864
11.1337 | -105.3715
11.1373 | -105.3568
11.141 | -105.3358
11.1447 | -105.317
11.1484 | -105.2984
11.152 | -105.2808
11.1557 | -105.2612
11.1594 | -105.2414
11.163 | -105.2172
11.1667 | -105.1906
11.1704 | -105.1693
11.1741 | -105.1497
11.1777 | -105.1264
11.1814 | -105.1042
11.1851 | -105.0845
11.1888 | -105.0672
11.1924 | -105.0521
11.1961 | -105.0424
11.1998 | -105.0316
11.2035 | -105.0182
11.2071 | -105.0073
11.2108 | -105.0059
11.2145 | -105.0075
11.2181 | -105.0087
11.2218 | -105.015
11.2255 | -105.0224
11.2292 | -105.032
11.2328 | -105.0432
11.2365 | -105.0512
11.2402 | -105.0656
11.2439 | -105.082
11.2475 | -105.0998
11.2512 | -105.1171
11.2549 | -105.1344
11.2586 | -105.1518
11.2622 | -105.1702
11.2659 | -105.1878
11.2696 | -105.2106
11.2732 | -105.229
11.2769 | -105.2498
11.2806 | -105.2741
11.2843 | -105.2951
11.2879 | -105.3166
11.2916 | -105.3385
11.2953 | -105.3639
11.299 | -105.3917
11.3026 | -105.4214
11.3063 | -105.4521
11.31 | -105.481
11.3137 | -105.5089
11.3173 | -105.5335
11.321 | -105.5566
11.3247 | -105.5778
11.3283 | -105.6004
11.332 | -105.6231
11.3357 | -105.6384
11.3394 | -105.6519
11.343 | -105.661
11.3467 | -105.6621
11.3504 | -105.6656
11.3541 | -105.6656
11.3577 | -105.662
11.3614 | -105.6537
11.3651 | -105.6435
11.3688 | -105.6365
11.3724 | -105.6261
11.3761 | -105.6146
11.3798 | -105.6014
11.3834 | -105.5851
11.3871 | -105.5647
11.3908 | -105.5488
11.3945 | -105.5351
11.3981 | -105.5186
11.4018 | -105.5029
11.4055 | -105.4831
11.4092 | -105.4645
11.4128 | -105.448
11.4165 | -105.4329
11.4202 | -105.4203
11.4238 | -105.4074
11.4275 | -105.3966
11.4312 | -105.3879
11.4349 | -105.3765
11.4385 | -105.3638
11.4422 | -105.3503
11.4459 | -105.3434
11.4496 | -105.3417
11.4532 | -105.3363
11.4569 | -105.3323
11.4606 | -105.3298
11.4643 | -105.3198
11.4679 | -105.3104
11.4716 | -105.301
11.4753 | -105.2937
11.4789 | -105.2895
11.4826 | -105.2885
11.4863 | -105.2903
11.49 | -105.2875
11.4936 | -105.2815
11.4973 | -105.2726
11.501 | -105.2637
11.5047 | -105.2553
11.5083 | -105.2513
11.512 | -105.2503
11.5157 | -105.2389
11.5194 | -105.2285
11.523 | -105.2189
11.5267 | -105.2122
11.5304 | -105.2054
11.534 | -105.1968
11.5377 | -105.1874
11.5414 | -105.1773
11.5451 | -105.172
11.5487 | -105.1676
11.5524 | -105.1615
11.5561 | -105.155
11.5598 | -105.1524
11.5634 | -105.155
11.5671 | -105.158
11.5708 | -105.1591
11.5745 | -105.1603
11.5781 | -105.1606
11.5818 | -105.1621
11.5855 | -105.1661
11.5891 | -105.1731
11.5928 | -105.1769
11.5965 | -105.1806
11.6002 | -105.1853
11.6038 | -105.1929
11.6075 | -105.1971
11.6112 | -105.1926
11.6149 | -105.1878
11.6185 | -105.1874
11.6222 | -105.1829
11.6259 | -105.1731
11.6296 | -105.1667
11.6332 | -105.1592
11.6369 | -105.1515
11.6406 | -105.1392
11.6442 | -105.1291
11.6479 | -105.12
11.6516 | -105.1094
11.6553 | -105.1109
11.6589 | -105.1102
11.6626 | -105.1141
11.6663 | -105.1208
11.67 | -105.1305
11.6736 | -105.1408
11.6773 | -105.1527
11.681 | -105.1679
11.6846 | -105.1915
11.6883 | -105.2132
11.692 | -105.2334
11.6957 | -105.2532
11.6993 | -105.2754
11.703 | -105.2972
11.7067 | -105.3183
11.7104 | -105.3419
11.714 | -105.3637
11.7177 | -105.3903
11.7214 | -105.4167
11.7251 | -105.4364
11.7287 | -105.4499
11.7324 | -105.462
11.7361 | -105.4749
11.7397 | -105.4876
11.7434 | -105.4992
11.7471 | -105.5084
11.7508 | -105.5184
11.7544 | -105.5241
11.7581 | -105.5274
11.7618 | -105.5329
11.7655 | -105.5361
11.7691 | -105.5375
11.7728 | -105.5401
11.7765 | -105.5428
11.7802 | -105.5475
11.7838 | -105.5504
11.7875 | -105.5502
11.7912 | -105.5481
11.7948 | -105.5494
11.7985 | -105.549
11.8022 | -105.5494
11.8059 | -105.5487
11.8095 | -105.547
11.8132 | -105.5443
11.8169 | -105.544
11.8206 | -105.544
11.8242 | -105.5435
11.8279 | -105.5447
11.8316 | -105.5435
11.8353 | -105.5406
11.8389 | -105.5365
11.8426 | -105.5332
11.8463 | -105.5334
11.8499 | -105.534
11.8536 | -105.5335
11.8573 | -105.5311
11.861 | -105.5314
11.8646 | -105.5341
11.8683 | -105.5329
11.872 | -105.5291
11.8757 | -105.5283
11.8793 | -105.5287
11.883 | -105.5241
11.8867 | -105.5179
11.8904 | -105.5101
11.894 | -105.5052
11.8977 | -105.4975
11.9014 | -105.4899
11.905 | -105.4845
11.9087 | -105.4795
11.9124 | -105.4763
11.9161 | -105.4737
11.9197 | -105.4732
11.9234 | -105.4705
11.9271 | -105.4723
11.9308 | -105.4768
11.9344 | -105.479
11.9381 | -105.4824
11.9418 | -105.4879
11.9455 | -105.4967
11.9491 | -105.5037
11.9528 | -105.5112
11.9565 | -105.5206
11.9601 | -105.5329
11.9638 | -105.5513
11.9675 | -105.5664
11.9712 | -105.581
11.9748 | -105.5951
11.9785 | -105.6094
11.9822 | -105.6283
11.9859 | -105.6466
11.9895 | -105.6651
11.9932 | -105.682
11.9969 | -105.7006
12.0006 | -105.7188
12.0042 | -105.7381
12.0079 | -105.7588
12.0116 | -105.7791
12.0152 | -105.7986
12.0189 | -105.817
12.0226 | -105.8383
12.0263 | -105.8582
12.0299 | -105.8757
12.0336 | -105.8901
12.0373 | -105.9022
12.041 | -105.9167
12.0446 | -105.927
12.0483 | -105.9324
12.052 | -105.9354
12.0556 | -105.9372
12.0593 | -105.9369
12.063 | -105.9322
12.0667 | -105.9227
12.0703 | -105.9163
12.074 | -105.9039
12.0777 | -105.8871
12.0814 | -105.8713
12.085 | -105.8514
12.0887 | -105.8303
12.0924 | -105.8139
12.0961 | -105.8003
12.0997 | -105.7852
12.1034 | -105.7688
12.1071 | -105.7573
12.1107 | -105.7505
12.1144 | -105.7453
12.1181 | -105.7399
12.1218 | -105.7401
12.1254 | -105.7424
12.1291 | -105.7455
12.1328 | -105.7473
12.1365 | -105.7489
12.1401 | -105.7467
12.1438 | -105.7442
12.1475 | -105.7431
12.1512 | -105.745
12.1548 | -105.7471
12.1585 | -105.7419
12.1622 | -105.7335
12.1658 | -105.7269
12.1695 | -105.7142
12.1732 | -105.6997
12.1769 | -105.6822
12.1805 | -105.6628
12.1842 | -105.6388
12.1879 | -105.6069
12.1916 | -105.584
12.1952 | -105.5594
12.1989 | -105.5333
12.2026 | -105.5101
12.2063 | -105.4893
12.2099 | -105.4694
12.2136 | -105.4474
12.2173 | -105.4252
12.2209 | -105.4089
12.2246 | -105.3947
12.2283 | -105.386
12.232 | -105.3786
12.2356 | -105.3735
12.2393 | -105.3734
12.243 | -105.3797
12.2467 | -105.3877
12.2503 | -105.4006
12.254 | -105.4157
12.2577 | -105.429
12.2614 | -105.4453
12.265 | -105.4631
12.2687 | -105.4779
12.2724 | -105.4945
12.276 | -105.5119
12.2797 | -105.5296
12.2834 | -105.5434
12.2871 | -105.553
12.2907 | -105.5663
12.2944 | -105.5784
12.2981 | -105.5915
12.3018 | -105.6036
12.3054 | -105.6138
12.3091 | -105.6273
12.3128 | -105.6391
12.3165 | -105.6521
12.3201 | -105.6635
12.3238 | -105.6687
12.3275 | -105.6684
12.3311 | -105.6678
12.3348 | -105.6727
12.3385 | -105.6744
12.3422 | -105.6736
12.3458 | -105.6744
12.3495 | -105.6718
12.3532 | -105.6655
12.3569 | -105.6572
12.3605 | -105.653
12.3642 | -105.6464
12.3679 | -105.6359
12.3716 | -105.6236
12.3752 | -105.6119
12.3789 | -105.6007
12.3826 | -105.5924
12.3862 | -105.5857
12.3899 | -105.5815
12.3936 | -105.5795
12.3973 | -105.5818
12.4009 | -105.5828
12.4046 | -105.584
12.4083 | -105.5829
12.412 | -105.5851
12.4156 | -105.5928
12.4193 | -105.6031
12.423 | -105.6108
12.4267 | -105.6198
12.4303 | -105.6314
12.434 | -105.642
12.4377 | -105.6582
12.4413 | -105.675
12.445 | -105.6889
12.4487 | -105.706
12.4524 | -105.7195
12.456 | -105.7333
12.4597 | -105.7446
12.4634 | -105.7571
12.4671 | -105.7694
12.4707 | -105.7748
12.4744 | -105.7778
12.4781 | -105.7799
12.4817 | -105.7827
12.4854 | -105.7874
12.4891 | -105.7864
12.4928 | -105.7864
12.4964 | -105.7864
12.5001 | -105.786
12.5038 | -105.784
12.5075 | -105.7846
12.5111 | -105.783
12.5148 | -105.7783
12.5185 | -105.775
12.5222 | -105.7709
12.5258 | -105.7663
12.5295 | -105.7626
12.5332 | -105.765
12.5368 | -105.7675
12.5405 | -105.7666
12.5442 | -105.7672
12.5479 | -105.7694
12.5515 | -105.7735
12.5552 | -105.7739
12.5589 | -105.7726
12.5626 | -105.771
12.5662 | -105.767
12.5699 | -105.7648
12.5736 | -105.7637
12.5773 | -105.757
12.5809 | -105.7519
12.5846 | -105.7496
12.5883 | -105.7472
12.5919 | -105.7449
12.5956 | -105.7379
12.5993 | -105.7331
12.603 | -105.7307
12.6066 | -105.7285
12.6103 | -105.7268
12.614 | -105.7289
12.6177 | -105.7298
12.6213 | -105.7288
12.625 | -105.7333
12.6287 | -105.7366
12.6324 | -105.742
12.636 | -105.749
12.6397 | -105.7609
12.6434 | -105.7716
12.647 | -105.777
12.6507 | -105.7842
12.6544 | -105.7943
12.6581 | -105.8023
12.6617 | -105.8087
12.6654 | -105.8109
12.6691 | -105.8144
12.6728 | -105.816
12.6764 | -105.8199
12.6801 | -105.8207
12.6838 | -105.8245
12.6875 | -105.8268
12.6911 | -105.8279
12.6948 | -105.8337
12.6985 | -105.8391
12.7021 | -105.8471
12.7058 | -105.8563
12.7095 | -105.8653
12.7132 | -105.869
12.7168 | -105.8787
12.7205 | -105.8857
12.7242 | -105.8908
12.7279 | -105.8984
12.7315 | -105.9073
12.7352 | -105.914
12.7389 | -105.9203
12.7425 | -105.9249
12.7462 | -105.9269
12.7499 | -105.9294
12.7536 | -105.9268
12.7572 | -105.9221
12.7609 | -105.919
12.7646 | -105.9118
12.7683 | -105.904
12.7719 | -105.8943
12.7756 | -105.8804
12.7793 | -105.8669
12.783 | -105.8526
12.7866 | -105.837
12.7903 | -105.8219
12.794 | -105.8074
12.7976 | -105.7921
12.8013 | -105.7799
12.805 | -105.7718
12.8087 | -105.7618
12.8123 | -105.7541
12.816 | -105.7487
12.8197 | -105.7455
12.8234 | -105.7438
12.827 | -105.7409
12.8307 | -105.7419
12.8344 | -105.7443
12.8381 | -105.7478
12.8417 | -105.7521
12.8454 | -105.7568
12.8491 | -105.7652
12.8527 | -105.7742
12.8564 | -105.7838
12.8601 | -105.7897
12.8638 | -105.7955
12.8674 | -105.7996
12.8711 | -105.7999
12.8748 | -105.8002
12.8785 | -105.7988
12.8821 | -105.7981
12.8858 | -105.7941
12.8895 | -105.7871
12.8932 | -105.7811
12.8968 | -105.7759
12.9005 | -105.7698
12.9042 | -105.7581
12.9078 | -105.7423
12.9115 | -105.7303
12.9152 | -105.718
12.9189 | -105.7066
12.9225 | -105.6982
12.9262 | -105.6924
12.9299 | -105.6871
12.9336 | -105.6809
12.9372 | -105.6773
12.9409 | -105.677
12.9446 | -105.6766
12.9483 | -105.6788
12.9519 | -105.6841
12.9556 | -105.6897
12.9593 | -105.701
12.9629 | -105.7158
12.9666 | -105.7298
12.9703 | -105.7412
12.974 | -105.7546
12.9776 | -105.7711
12.9813 | -105.7874
12.985 | -105.804
12.9887 | -105.818
12.9923 | -105.8303
12.996 | -105.8459
12.9997 | -105.8576
13.0034 | -105.8683
13.007 | -105.8799
13.0107 | -105.89
13.0144 | -105.9046
13.018 | -105.9172
13.0217 | -105.9272
13.0254 | -105.9351
13.0291 | -105.9428
13.0327 | -105.9496
13.0364 | -105.954
13.0401 | -105.9585
13.0438 | -105.9621
13.0474 | -105.9657
13.0511 | -105.9679
13.0548 | -105.968
13.0585 | -105.9711
13.0621 | -105.9714
13.0658 | -105.9695
13.0695 | -105.9639
13.0731 | -105.9585
13.0768 | -105.9495
13.0805 | -105.9369
13.0842 | -105.9212
13.0878 | -105.9031
13.0915 | -105.8856
13.0952 | -105.8703
13.0989 | -105.8517
13.1025 | -105.827
13.1062 | -105.8003
13.1099 | -105.773
13.1136 | -105.7448
13.1172 | -105.7198
13.1209 | -105.6962
13.1246 | -105.6697
13.1282 | -105.6444
13.1319 | -105.6221
13.1356 | -105.5984
13.1393 | -105.5751
13.1429 | -105.5532
13.1466 | -105.5301
13.1503 | -105.5095
13.154 | -105.4934
13.1576 | -105.4811
13.1613 | -105.4701
13.165 | -105.4592
13.1686 | -105.4503
13.1723 | -105.4417
13.176 | -105.4386
13.1797 | -105.4381
13.1833 | -105.4347
13.187 | -105.4331
13.1907 | -105.4316
13.1944 | -105.4308
13.198 | -105.4282
13.2017 | -105.4262
13.2054 | -105.4255
13.2091 | -105.4257
13.2127 | -105.4259
13.2164 | -105.4261
13.2201 | -105.4281
13.2237 | -105.4318
13.2274 | -105.4349
13.2311 | -105.4371
13.2348 | -105.4398
13.2384 | -105.4421
13.2421 | -105.4453
13.2458 | -105.4522
13.2495 | -105.4562
13.2531 | -105.459
13.2568 | -105.4635
13.2605 | -105.4673
13.2642 | -105.4709
13.2678 | -105.476
13.2715 | -105.4823
13.2752 | -105.4852
13.2788 | -105.486
13.2825 | -105.4848
13.2862 | -105.4814
13.2899 | -105.4794
13.2935 | -105.4777
13.2972 | -105.4742
13.3009 | -105.4707
13.3046 | -105.4659
13.3082 | -105.4575
13.3119 | -105.448
13.3156 | -105.4377
13.3193 | -105.4266
13.3229 | -105.4137
13.3266 | -105.4
13.3303 | -105.3858
13.3339 | -105.3702
13.3376 | -105.3548
13.3413 | -105.339
13.345 | -105.3213
13.3486 | -105.3021
13.3523 | -105.2814
13.356 | -105.2624
13.3597 | -105.2428
13.3633 | -105.2254
13.367 | -105.2062
13.3707 | -105.1896
13.3744 | -105.1776
13.378 | -105.1678
13.3817 | -105.1603
13.3854 | -105.1513
13.389 | -105.148
13.3927 | -105.1446
13.3964 | -105.1444
13.4001 | -105.146
13.4037 | -105.1513
13.4074 | -105.1594
13.4111 | -105.1685
13.4148 | -105.1805
13.4184 | -105.1954
13.4221 | -105.2124
13.4258 | -105.2341
13.4294 | -105.2591
13.4331 | -105.2856
13.4368 | -105.3121
13.4405 | -105.3402
13.4441 | -105.3703
13.4478 | -105.4021
13.4515 | -105.4319
13.4552 | -105.4641
13.4588 | -105.497
13.4625 | -105.5299
13.4662 | -105.56
13.4699 | -105.5929
13.4735 | -105.6234
13.4772 | -105.6513
13.4809 | -105.677
13.4845 | -105.6966
13.4882 | -105.7148
13.4919 | -105.7325
13.4956 | -105.7472
13.4992 | -105.7619
13.5029 | -105.7757
13.5066 | -105.7868
13.5103 | -105.7986
13.5139 | -105.8071
13.5176 | -105.8151
13.5213 | -105.8201
13.525 | -105.8226
13.5286 | -105.8267
13.5323 | -105.8302
13.536 | -105.8331
13.5396 | -105.8384
13.5433 | -105.8459
13.547 | -105.8552
13.5507 | -105.8613
13.5543 | -105.8696
13.558 | -105.8805
13.5617 | -105.8903
13.5654 | -105.8976
13.569 | -105.9041
13.5727 | -105.9137
13.5764 | -105.9237
13.5801 | -105.936
13.5837 | -105.946
13.5874 | -105.9591
13.5911 | -105.9707
13.5947 | -105.9783
13.5984 | -105.9826
13.6021 | -105.9861
13.6058 | -105.9898
13.6094 | -105.9904
13.6131 | -105.9873
13.6168 | -105.9826
13.6205 | -105.9747
13.6241 | -105.9674
13.6278 | -105.9581
13.6315 | -105.9461
13.6352 | -105.931
13.6388 | -105.9151
13.6425 | -105.9008
13.6462 | -105.8853
13.6498 | -105.8697
13.6535 | -105.8547
13.6572 | -105.839
13.6609 | -105.824
13.6645 | -105.81
13.6682 | -105.8002
13.6719 | -105.791
13.6756 | -105.7829
13.6792 | -105.7747
13.6829 | -105.7681
13.6866 | -105.7618
13.6903 | -105.7563
13.6939 | -105.752
13.6976 | -105.7499
13.7013 | -105.7485
13.7049 | -105.7457
13.7086 | -105.7463
13.7123 | -105.7477
13.716 | -105.7506
13.7196 | -105.7534
13.7233 | -105.7559
13.727 | -105.7568
13.7307 | -105.7603
13.7343 | -105.7654
13.738 | -105.769
13.7417 | -105.7736
13.7454 | -105.7804
13.749 | -105.7864
13.7527 | -105.7913
13.7564 | -105.7989
13.76 | -105.8072
13.7637 | -105.811
13.7674 | -105.8124
13.7711 | -105.8126
13.7747 | -105.8148
13.7784 | -105.8163
13.7821 | -105.8194
13.7858 | -105.8201
13.7894 | -105.8201
13.7931 | -105.822
13.7968 | -105.8262
13.8004 | -105.8298
13.8041 | -105.8329
13.8078 | -105.8378
13.8115 | -105.8467
13.8151 | -105.8568
13.8188 | -105.8656
13.8225 | -105.8722
13.8262 | -105.8823
13.8298 | -105.8981
13.8335 | -105.9157
13.8372 | -105.935
13.8409 | -105.9509
13.8445 | -105.966
13.8482 | -105.9844
13.8519 | -106.0044
13.8555 | -106.0251
13.8592 | -106.0459
13.8629 | -106.0663
13.8666 | -106.0845
13.8702 | -106.1017
13.8739 | -106.118
13.8776 | -106.1332
13.8813 | -106.1459
13.8849 | -106.1556
13.8886 | -106.1627
13.8923 | -106.1647
13.896 | -106.1633
13.8996 | -106.166
13.9033 | -106.1666
13.907 | -106.1579
13.9106 | -106.1464
13.9143 | -106.1339
13.918 | -106.1186
13.9217 | -106.0986
13.9253 | -106.0803
13.929 | -106.0559
13.9327 | -106.0267
13.9364 | -106.0005
13.94 | -105.9772
13.9437 | -105.9496
13.9474 | -105.9148
13.9511 | -105.8796
13.9547 | -105.846
13.9584 | -105.8091
13.9621 | -105.7725
13.9657 | -105.7325
13.9694 | -105.6918
13.9731 | -105.6509
13.9768 | -105.615
13.9804 | -105.5792
13.9841 | -105.5434
13.9878 | -105.5071
13.9915 | -105.4749
13.9951 | -105.4469
13.9988 | -105.4184
14.0025 | -105.3923
14.0062 | -105.3676
14.0098 | -105.3443
14.0135 | -105.3251
14.0172 | -105.3084
14.0208 | -105.2957
14.0245 | -105.2864
14.0282 | -105.2776
14.0319 | -105.2724
14.0355 | -105.269
14.0392 | -105.2663
14.0429 | -105.2702
14.0466 | -105.2803
14.0502 | -105.2893
14.0539 | -105.2974
14.0576 | -105.3117
14.0613 | -105.33
14.0649 | -105.3506
14.0686 | -105.3741
14.0723 | -105.3998
14.0759 | -105.427
14.0796 | -105.4587
14.0833 | -105.4927
14.087 | -105.5291
14.0906 | -105.5663
14.0943 | -105.6034
14.098 | -105.6371
14.1017 | -105.6685
14.1053 | -105.7003
14.109 | -105.7365
14.1127 | -105.768
14.1164 | -105.7993
14.12 | -105.8299
14.1237 | -105.8584
14.1274 | -105.8796
14.131 | -105.8993
14.1347 | -105.9163
14.1384 | -105.9301
14.1421 | -105.9377
14.1457 | -105.9456
14.1494 | -105.9573
14.1531 | -105.9671
14.1568 | -105.9715
14.1604 | -105.979
14.1641 | -105.9841
14.1678 | -105.9901
14.1715 | -105.9937
14.1751 | -105.9886
14.1788 | -105.9779
14.1825 | -105.9703
14.1861 | -105.9644
14.1898 | -105.9618
14.1935 | -105.956
14.1972 | -105.9467
14.2008 | -105.939
14.2045 | -105.9317
14.2082 | -105.9249
14.2119 | -105.9191
14.2155 | -105.9115
14.2192 | -105.906
14.2229 | -105.9037
14.2265 | -105.9016
14.2302 | -105.9003
14.2339 | -105.8985
14.2376 | -105.8994
14.2412 | -105.8999
14.2449 | -105.9018
14.2486 | -105.9036
14.2523 | -105.9074
14.2559 | -105.9157
14.2596 | -105.9246
14.2633 | -105.9337
14.267 | -105.9416
14.2706 | -105.951
14.2743 | -105.962
14.278 | -105.9737
14.2816 | -105.9847
14.2853 | -105.9948
14.289 | -106.0052
14.2927 | -106.0164
14.2963 | -106.0327
14.3 | -106.0517
14.3037 | -106.0645
14.3074 | -106.0729
14.311 | -106.0826
14.3147 | -106.0889
14.3184 | -106.0951
14.3221 | -106.0998
14.3257 | -106.1014
14.3294 | -106.1035
14.3331 | -106.1057
14.3367 | -106.1072
14.3404 | -106.1065
14.3441 | -106.1026
14.3478 | -106.0947
14.3514 | -106.0866
14.3551 | -106.0764
14.3588 | -106.0638
14.3625 | -106.0516
14.3661 | -106.0401
14.3698 | -106.0259
14.3735 | -106.0087
14.3772 | -105.9949
14.3808 | -105.9844
14.3845 | -105.9719
14.3882 | -105.956
14.3918 | -105.9393
14.3955 | -105.9205
14.3992 | -105.9067
14.4029 | -105.8981
14.4065 | -105.8894
14.4102 | -105.8812
14.4139 | -105.876
14.4176 | -105.8702
14.4212 | -105.8661
14.4249 | -105.8665
14.4286 | -105.8643
14.4323 | -105.861
14.4359 | -105.8622
14.4396 | -105.8656
14.4433 | -105.8715
14.4469 | -105.8809
14.4506 | -105.8863
14.4543 | -105.89
14.458 | -105.8954
14.4616 | -105.9021
14.4653 | -105.9088
14.469 | -105.9151
14.4727 | -105.9237
14.4763 | -105.9311
14.48 | -105.9413
14.4837 | -105.9493
14.4873 | -105.9544
14.491 | -105.9616
14.4947 | -105.9677
14.4984 | -105.9727
14.502 | -105.9756
14.5057 | -105.9774
14.5094 | -105.9834
14.5131 | -105.9887
14.5167 | -105.9926
14.5204 | -105.999
14.5241 | -106.0082
14.5278 | -106.0174
14.5314 | -106.0213
14.5351 | -106.0263
14.5388 | -106.0296
14.5424 | -106.0341
14.5461 | -106.0372
14.5498 | -106.0402
14.5535 | -106.0451
14.5571 | -106.0507
14.5608 | -106.0531
14.5645 | -106.0545
14.5682 | -106.0555
14.5718 | -106.0546
14.5755 | -106.0515
14.5792 | -106.05
14.5829 | -106.0501
14.5865 | -106.0462
14.5902 | -106.0421
14.5939 | -106.0389
14.5975 | -106.0342
14.6012 | -106.0272
14.6049 | -106.0226
14.6086 | -106.0136
14.6122 | -106.003
14.6159 | -105.9963
14.6196 | -105.9912
14.6233 | -105.9837
14.6269 | -105.9768
14.6306 | -105.9706
14.6343 | -105.9628
14.638 | -105.9575
14.6416 | -105.9546
14.6453 | -105.9522
14.649 | -105.9482
14.6526 | -105.9421
14.6563 | -105.9363
14.66 | -105.9319
14.6637 | -105.9287
14.6673 | -105.9275
14.671 | -105.9247
14.6747 | -105.9203
14.6784 | -105.9187
14.682 | -105.9167
14.6857 | -105.9146
14.6894 | -105.909
14.6931 | -105.9077
14.6967 | -105.9082
14.7004 | -105.9061
14.7041 | -105.9032
14.7077 | -105.9023
14.7114 | -105.9
14.7151 | -105.8993
14.7188 | -105.8975
14.7224 | -105.8969
14.7261 | -105.8945
14.7298 | -105.8855
14.7335 | -105.8778
14.7371 | -105.8721
14.7408 | -105.8678
14.7445 | -105.8633
14.7482 | -105.8586
14.7518 | -105.851
14.7555 | -105.8459
14.7592 | -105.8404
14.7628 | -105.8314
14.7665 | -105.8282
14.7702 | -105.8235
14.7739 | -105.8182
14.7775 | -105.8152
14.7812 | -105.8129
14.7849 | -105.8098
14.7886 | -105.8092
14.7922 | -105.808
14.7959 | -105.8074
14.7996 | -105.8059
14.8033 | -105.8042
14.8069 | -105.8054
14.8106 | -105.8079
14.8143 | -105.8068
14.8179 | -105.8067
14.8216 | -105.8072
14.8253 | -105.8071
14.829 | -105.8065
14.8326 | -105.8099
14.8363 | -105.8162
14.84 | -105.8218
14.8437 | -105.8295
14.8473 | -105.8367
14.851 | -105.8422
14.8547 | -105.8504
14.8583 | -105.8588
14.862 | -105.8686
14.8657 | -105.8829
14.8694 | -105.8981
14.873 | -105.9142
14.8767 | -105.9311
14.8804 | -105.9425
14.8841 | -105.9521
14.8877 | -105.9655
14.8914 | -105.9787
14.8951 | -105.9936
14.8988 | -106.0098
14.9024 | -106.0261
14.9061 | -106.0396
14.9098 | -106.0517
14.9134 | -106.0636
14.9171 | -106.0755
14.9208 | -106.082
14.9245 | -106.0904
14.9281 | -106.0962
14.9318 | -106.1013
14.9355 | -106.1066
14.9392 | -106.1144
14.9428 | -106.1217
14.9465 | -106.1214
14.9502 | -106.1185
14.9539 | -106.116
14.9575 | -106.1125
14.9612 | -106.1051
14.9649 | -106.0982
14.9685 | -106.0901
14.9722 | -106.0818
14.9759 | -106.0744
14.9796 | -106.063
14.9832 | -106.0515
14.9869 | -106.038
14.9906 | -106.0203
14.9943 | -106.0058
14.9979 | -105.993
15.0016 | -105.9796
15.0053 | -105.9674
15.009 | -105.9531
15.0126 | -105.9393
15.0163 | -105.9268
15.02 | -105.913
15.0236 | -105.9003
15.0273 | -105.8874
15.031 | -105.8744
15.0347 | -105.8641
15.0383 | -105.8527
15.042 | -105.8417
15.0457 | -105.8295
15.0494 | -105.8201
15.053 | -105.809
15.0567 | -105.798
15.0604 | -105.7857
15.0641 | -105.7757
15.0677 | -105.7684
15.0714 | -105.7624
15.0751 | -105.7548
15.0787 | -105.7508
15.0824 | -105.746
15.0861 | -105.7438
15.0898 | -105.7438
15.0934 | -105.7447
15.0971 | -105.7506
15.1008 | -105.7613
15.1045 | -105.7701
15.1081 | -105.7783
15.1118 | -105.7847
15.1155 | -105.791
15.1192 | -105.8011
15.1228 | -105.8132
15.1265 | -105.8245
15.1302 | -105.8327
15.1338 | -105.8425
15.1375 | -105.8532
15.1412 | -105.8607
15.1449 | -105.8662
15.1485 | -105.8678
15.1522 | -105.8674
15.1559 | -105.8664
15.1596 | -105.8678
15.1632 | -105.8668
15.1669 | -105.8641
15.1706 | -105.8592
15.1743 | -105.8496
15.1779 | -105.8421
15.1816 | -105.8318
15.1853 | -105.8178
15.1889 | -105.8042
15.1926 | -105.7895
15.1963 | -105.7725
15.2 | -105.7587
15.2036 | -105.7463
15.2073 | -105.7343
15.211 | -105.7184
15.2147 | -105.7031
15.2183 | -105.6903
15.222 | -105.6776
15.2257 | -105.6666
15.2293 | -105.6573
15.233 | -105.6496
15.2367 | -105.6432
15.2404 | -105.6354
15.244 | -105.631
15.2477 | -105.6306
15.2514 | -105.6281
15.2551 | -105.6283
15.2587 | -105.6318
15.2624 | -105.6352
15.2661 | -105.6388
15.2698 | -105.6475
15.2734 | -105.6576
15.2771 | -105.6701
15.2808 | -105.6842
15.2844 | -105.6988
15.2881 | -105.7127
15.2918 | -105.731
15.2955 | -105.7489
15.2991 | -105.769
15.3028 | -105.7889
15.3065 | -105.8082
15.3102 | -105.8285
15.3138 | -105.8501
15.3175 | -105.8706
15.3212 | -105.8914
15.3249 | -105.912
15.3285 | -105.9345
15.3322 | -105.9585
15.3359 | -105.9781
15.3395 | -105.9951
15.3432 | -106.0134
15.3469 | -106.0346
15.3506 | -106.0531
15.3542 | -106.0682
15.3579 | -106.0802
15.3616 | -106.0938
15.3653 | -106.1013
15.3689 | -106.115
15.3726 | -106.1256
15.3763 | -106.1339
15.38 | -106.142
15.3836 | -106.1509
15.3873 | -106.1571
15.391 | -106.1619
15.3946 | -106.1653
15.3983 | -106.1678
15.402 | -106.1727
15.4057 | -106.1781
15.4093 | -106.1822
15.413 | -106.1844
15.4167 | -106.1842
15.4204 | -106.182
15.424 | -106.1778
15.4277 | -106.1733
15.4314 | -106.1704
15.4351 | -106.1657
15.4387 | -106.1588
15.4424 | -106.1494
15.4461 | -106.141
15.4497 | -106.1334
15.4534 | -106.1234
15.4571 | -106.111
15.4608 | -106.0996
15.4644 | -106.0873
15.4681 | -106.074
15.4718 | -106.0596
15.4755 | -106.0431
15.4791 | -106.025
15.4828 | -106.0108
15.4865 | -105.9983
15.4902 | -105.9835
15.4938 | -105.9685
15.4975 | -105.9555
15.5012 | -105.9411
15.5048 | -105.9251
15.5085 | -105.9084
15.5122 | -105.8928
15.5159 | -105.8787
15.5195 | -105.8636
15.5232 | -105.8502
15.5269 | -105.8381
15.5306 | -105.8272
15.5342 | -105.816
15.5379 | -105.8065
15.5416 | -105.7989
15.5452 | -105.7922
15.5489 | -105.7866
15.5526 | -105.7791
15.5563 | -105.7723
15.5599 | -105.7664
15.5636 | -105.7617
15.5673 | -105.7559
15.571 | -105.7505
15.5746 | -105.7461
15.5783 | -105.7423
15.582 | -105.7366
15.5857 | -105.7298
15.5893 | -105.7256
15.593 | -105.7215
15.5967 | -105.7152
15.6003 | -105.7087
15.604 | -105.702
15.6077 | -105.697
15.6114 | -105.6927
15.615 | -105.6868
15.6187 | -105.6824
15.6224 | -105.6768
15.6261 | -105.6686
15.6297 | -105.661
15.6334 | -105.6538
15.6371 | -105.6483
15.6408 | -105.6457
15.6444 | -105.6442
15.6481 | -105.6431
15.6518 | -105.6418
15.6554 | -105.6397
15.6591 | -105.6377
15.6628 | -105.6387
15.6665 | -105.6412
15.6701 | -105.6457
15.6738 | -105.6503
15.6775 | -105.6561
15.6812 | -105.6629
15.6848 | -105.6678
15.6885 | -105.675
15.6922 | -105.6804
15.6959 | -105.6884
15.6995 | -105.6982
15.7032 | -105.7057
15.7069 | -105.7109
15.7105 | -105.7145
15.7142 | -105.7211
15.7179 | -105.7301
15.7216 | -105.7361
15.7252 | -105.7423
15.7289 | -105.7515
15.7326 | -105.7594
15.7363 | -105.766
15.7399 | -105.7711
15.7436 | -105.7773
15.7473 | -105.7843
15.751 | -105.791
15.7546 | -105.7961
15.7583 | -105.8009
15.762 | -105.8038
15.7656 | -105.8072
15.7693 | -105.811
15.773 | -105.8118
15.7767 | -105.8143
15.7803 | -105.8157
15.784 | -105.8158
15.7877 | -105.8153
15.7914 | -105.8149
15.795 | -105.8144
15.7987 | -105.8146
15.8024 | -105.8136
15.8061 | -105.8138
15.8097 | -105.8144
15.8134 | -105.8123
15.8171 | -105.8089
15.8207 | -105.8087
15.8244 | -105.8084
15.8281 | -105.8032
15.8318 | -105.798
15.8354 | -105.7939
15.8391 | -105.7891
15.8428 | -105.7839
15.8465 | -105.7758
15.8501 | -105.7684
15.8538 | -105.7596
15.8575 | -105.7503
15.8612 | -105.7409
15.8648 | -105.7298
15.8685 | -105.7173
15.8722 | -105.7043
15.8758 | -105.69
15.8795 | -105.6748
15.8832 | -105.6579
15.8869 | -105.6402
15.8905 | -105.6225
15.8942 | -105.6021
15.8979 | -105.5789
15.9016 | -105.5586
15.9052 | -105.5391
15.9089 | -105.5215
15.9126 | -105.5079
15.9162 | -105.495
15.9199 | -105.4822
15.9236 | -105.4699
15.9273 | -105.4589
15.9309 | -105.4502
15.9346 | -105.4451
15.9383 | -105.4395
15.942 | -105.4343
15.9456 | -105.4271
15.9493 | -105.4218
15.953 | -105.4201
15.9567 | -105.4205
15.9603 | -105.4224
15.964 | -105.4216
15.9677 | -105.4225
15.9713 | -105.4247
15.975 | -105.4245
15.9787 | -105.424
15.9824 | -105.423
15.986 | -105.4218
15.9897 | -105.4218
15.9934 | -105.4213
15.9971 | -105.4211
16.0007 | -105.421
16.0044 | -105.4164
16.0081 | -105.4141
16.0118 | -105.4126
16.0154 | -105.4084
16.0191 | -105.4031
16.0228 | -105.3938
16.0264 | -105.385
16.0301 | -105.3766
16.0338 | -105.3666
16.0375 | -105.3566
16.0411 | -105.3458
16.0448 | -105.3357
16.0485 | -105.3257
16.0522 | -105.3168
16.0558 | -105.3081
16.0595 | -105.2984
16.0632 | -105.289
16.0669 | -105.2818
16.0705 | -105.2739
16.0742 | -105.2653
16.0779 | -105.2556
16.0815 | -105.2477
16.0852 | -105.2408
16.0889 | -105.2341
16.0926 | -105.2298
16.0962 | -105.2263
16.0999 | -105.2203
16.1036 | -105.2164
16.1073 | -105.2143
16.1109 | -105.2134
16.1146 | -105.2098
16.1183 | -105.208
16.122 | -105.2081
16.1256 | -105.2115
16.1293 | -105.2161
16.133 | -105.2171
16.1366 | -105.2188
16.1403 | -105.2199
16.144 | -105.2188
16.1477 | -105.2173
16.1513 | -105.2178
16.155 | -105.2196
16.1587 | -105.2224
16.1624 | -105.2259
16.166 | -105.2278
16.1697 | -105.2311
16.1734 | -105.234
16.1771 | -105.237
16.1807 | -105.2405
16.1844 | -105.2438
16.1881 | -105.2444
16.1917 | -105.2457
16.1954 | -105.2506
16.1991 | -105.2562
16.2028 | -105.2622
16.2064 | -105.269
16.2101 | -105.277
16.2138 | -105.2855
16.2175 | -105.2949
16.2211 | -105.3053
16.2248 | -105.3166
16.2285 | -105.328
16.2322 | -105.3403
16.2358 | -105.3553
16.2395 | -105.3684
16.2432 | -105.3836
16.2468 | -105.3989
16.2505 | -105.4153
16.2542 | -105.4352
16.2579 | -105.4536
16.2615 | -105.4703
16.2652 | -105.4861
16.2689 | -105.5013
16.2726 | -105.5193
16.2762 | -105.5343
16.2799 | -105.5488
16.2836 | -105.565
16.2872 | -105.5823
16.2909 | -105.5989
16.2946 | -105.6143
16.2983 | -105.6288
16.3019 | -105.6435
16.3056 | -105.6593
16.3093 | -105.6752
16.313 | -105.6867
16.3166 | -105.6975
16.3203 | -105.7094
16.324 | -105.718
16.3277 | -105.725
16.3313 | -105.7354
16.335 | -105.7417
16.3387 | -105.7475
16.3423 | -105.7552
16.346 | -105.7623
16.3497 | -105.7678
16.3534 | -105.7726
16.357 | -105.7748
16.3607 | -105.7743
16.3644 | -105.7759
16.3681 | -105.7784
16.3717 | -105.7812
16.3754 | -105.7849
16.3791 | -105.7868
16.3828 | -105.7888
16.3864 | -105.7917
16.3901 | -105.7933
16.3938 | -105.7957
16.3974 | -105.8
16.4011 | -105.8057
16.4048 | -105.8105
16.4085 | -105.8208
16.4121 | -105.8323
16.4158 | -105.8409
16.4195 | -105.8513
16.4232 | -105.8634
16.4268 | -105.8791
16.4305 | -105.8914
16.4342 | -105.9008
16.4379 | -105.9118
16.4415 | -105.924
16.4452 | -105.9347
16.4489 | -105.9462
16.4525 | -105.9594
16.4562 | -105.9731
16.4599 | -105.984
16.4636 | -105.9932
16.4672 | -106.0035
16.4709 | -106.0144
16.4746 | -106.0233
16.4783 | -106.0298
16.4819 | -106.035
16.4856 | -106.0361
16.4893 | -106.0362
16.493 | -106.0359
16.4966 | -106.0351
16.5003 | -106.0349
16.504 | -106.0298
16.5076 | -106.0212
16.5113 | -106.0128
16.515 | -106.0015
16.5187 | -105.986
16.5223 | -105.9723
16.526 | -105.9577
16.5297 | -105.9415
16.5334 | -105.9265
16.537 | -105.9114
16.5407 | -105.8971
16.5444 | -105.8835
16.5481 | -105.8704
16.5517 | -105.8545
16.5554 | -105.8371
16.5591 | -105.825
16.5627 | -105.8124
16.5664 | -105.8013
16.5701 | -105.7925
16.5738 | -105.7839
16.5774 | -105.7723
16.5811 | -105.7628
16.5848 | -105.7559
16.5885 | -105.7484
16.5921 | -105.7398
16.5958 | -105.7363
16.5995 | -105.7334
16.6031 | -105.729
16.6068 | -105.7247
16.6105 | -105.7207
16.6142 | -105.7199
16.6178 | -105.7213
16.6215 | -105.7208
16.6252 | -105.7222
16.6289 | -105.7245
16.6325 | -105.7268
16.6362 | -105.7306
16.6399 | -105.7365
16.6436 | -105.7402
16.6472 | -105.7435
16.6509 | -105.7496
16.6546 | -105.7565
16.6582 | -105.7639
16.6619 | -105.7721
16.6656 | -105.7818
16.6693 | -105.7884
16.6729 | -105.7958
16.6766 | -105.8014
16.6803 | -105.8094
16.684 | -105.8201
16.6876 | -105.8314
16.6913 | -105.8396
16.695 | -105.8468
16.6987 | -105.8542
16.7023 | -105.8635
16.706 | -105.8704
16.7097 | -105.8762
16.7133 | -105.8816
16.717 | -105.8857
16.7207 | -105.8906
16.7244 | -105.8975
16.728 | -105.9025
16.7317 | -105.9054
16.7354 | -105.9074
16.7391 | -105.9101
16.7427 | -105.9151
16.7464 | -105.9175
16.7501 | -105.9208
16.7538 | -105.923
16.7574 | -105.9235
16.7611 | -105.9227
16.7648 | -105.9189
16.7684 | -105.9156
16.7721 | -105.9112
16.7758 | -105.9063
16.7795 | -105.904
16.7831 | -105.899
16.7868 | -105.8934
16.7905 | -105.8916
16.7942 | -105.8924
16.7978 | -105.8872
16.8015 | -105.8842
16.8052 | -105.879
16.8089 | -105.8748
16.8125 | -105.8687
16.8162 | -105.8635
16.8199 | -105.8601
16.8235 | -105.854
16.8272 | -105.8497
16.8309 | -105.8465
16.8346 | -105.8365
16.8382 | -105.8267
16.8419 | -105.8166
16.8456 | -105.8088
16.8493 | -105.8011
16.8529 | -105.7887
16.8566 | -105.7745
16.8603 | -105.7639
16.864 | -105.752
16.8676 | -105.7395
16.8713 | -105.7253
16.875 | -105.7123
16.8786 | -105.6978
16.8823 | -105.6837
16.886 | -105.6692
16.8897 | -105.6522
16.8933 | -105.6334
16.897 | -105.6125
16.9007 | -105.5919
16.9044 | -105.572
16.908 | -105.5554
16.9117 | -105.5377
16.9154 | -105.5212
16.919 | -105.4995
16.9227 | -105.478
16.9264 | -105.4582
16.9301 | -105.438
16.9337 | -105.4165
16.9374 | -105.3983
16.9411 | -105.3815
16.9448 | -105.367
16.9484 | -105.3507
16.9521 | -105.3347
16.9558 | -105.3249
16.9595 | -105.3121
16.9631 | -105.3004
16.9668 | -105.2904
16.9705 | -105.2814
16.9741 | -105.274
16.9778 | -105.2702
16.9815 | -105.2648
16.9852 | -105.2614
16.9888 | -105.2608
16.9925 | -105.2599
16.9962 | -105.2613
16.9999 | -105.2649
17.0035 | -105.27
17.0072 | -105.2779
17.0109 | -105.2862
17.0146 | -105.2967
17.0182 | -105.3076
17.0219 | -105.317
17.0256 | -105.3299
17.0292 | -105.3439
17.0329 | -105.3573
17.0366 | -105.371
17.0403 | -105.3882
17.0439 | -105.4052
17.0476 | -105.424
17.0513 | -105.4422
17.055 | -105.4604
17.0586 | -105.4781
17.0623 | -105.4944
17.066 | -105.5127
17.0697 | -105.5303
17.0733 | -105.5481
17.077 | -105.5693
17.0807 | -105.5865
17.0843 | -105.5994
17.088 | -105.6126
17.0917 | -105.6247
17.0954 | -105.6374
17.099 | -105.6479
17.1027 | -105.6557
17.1064 | -105.6647
17.1101 | -105.6711
17.1137 | -105.6728
17.1174 | -105.6755
17.1211 | -105.6786
17.1248 | -105.6819
17.1284 | -105.6847
17.1321 | -105.6856
17.1358 | -105.6857
17.1394 | -105.6864
17.1431 | -105.6844
17.1468 | -105.6799
17.1505 | -105.6738
17.1541 | -105.6689
17.1578 | -105.662
17.1615 | -105.6579
17.1652 | -105.6518
17.1688 | -105.6432
17.1725 | -105.6388
17.1762 | -105.6339
17.1799 | -105.6268
17.1835 | -105.6186
17.1872 | -105.61
17.1909 | -105.602
17.1945 | -105.5968
17.1982 | -105.5901
17.2019 | -105.5813
17.2056 | -105.5727
17.2092 | -105.5616
17.2129 | -105.5483
17.2166 | -105.5357
17.2203 | -105.5252
17.2239 | -105.513
17.2276 | -105.5014
17.2313 | -105.4903
17.235 | -105.4809
17.2386 | -105.4709
17.2423 | -105.4608
17.246 | -105.4498
17.2496 | -105.4395
17.2533 | -105.4288
17.257 | -105.4199
17.2607 | -105.4088
17.2643 | -105.3976
17.268 | -105.3876
17.2717 | -105.3791
17.2754 | -105.3706
17.279 | -105.3629
17.2827 | -105.3541
17.2864 | -105.3445
17.2901 | -105.3376
17.2937 | -105.3326
17.2974 | -105.3299
17.3011 | -105.3269
17.3047 | -105.3231
17.3084 | -105.3185
17.3121 | -105.3135
17.3158 | -105.3113
17.3194 | -105.3123
17.3231 | -105.312
17.3268 | -105.3111
17.3305 | -105.3147
17.3341 | -105.3164
17.3378 | -105.3168
17.3415 | -105.3177
17.3451 | -105.3197
17.3488 | -105.3222
17.3525 | -105.3248
17.3562 | -105.3278
17.3598 | -105.3352
17.3635 | -105.3414
17.3672 | -105.3471
17.3709 | -105.3497
17.3745 | -105.3541
17.3782 | -105.3613
17.3819 | -105.3701
17.3856 | -105.3789
17.3892 | -105.3889
17.3929 | -105.3995
17.3966 | -105.41
17.4002 | -105.4186
17.4039 | -105.4269
17.4076 | -105.4338
17.4113 | -105.4387
17.4149 | -105.4446
17.4186 | -105.4481
17.4223 | -105.4527
17.426 | -105.4559
17.4296 | -105.4585
17.4333 | -105.4594
17.437 | -105.4576
17.4407 | -105.4519
17.4443 | -105.4464
17.448 | -105.438
17.4517 | -105.4285
17.4553 | -105.4197
17.459 | -105.4112
17.4627 | -105.3997
17.4664 | -105.3847
17.47 | -105.3665
17.4737 | -105.3471
17.4774 | -105.3235
17.4811 | -105.2985
17.4847 | -105.2661
17.4884 | -105.2414
17.4921 | -105.2174
17.4958 | -105.1928
17.4994 | -105.1646
17.5031 | -105.1377
17.5068 | -105.1096
17.5104 | -105.0824
17.5141 | -105.0555
17.5178 | -105.0288
17.5215 | -105.0017
17.5251 | -104.9761
17.5288 | -104.9488
17.5325 | -104.9226
17.5362 | -104.8961
17.5398 | -104.8679
17.5435 | -104.8399
17.5472 | -104.8118
17.5509 | -104.7869
17.5545 | -104.7633
17.5582 | -104.7415
17.5619 | -104.7219
17.5655 | -104.7
17.5692 | -104.678
17.5729 | -104.6596
17.5766 | -104.6455
17.5802 | -104.6314
17.5839 | -104.6159
17.5876 | -104.6023
17.5913 | -104.5921
17.5949 | -104.5836
17.5986 | -104.5752
17.6023 | -104.5669
17.606 | -104.5623
17.6096 | -104.5595
17.6133 | -104.5548
17.617 | -104.553
17.6206 | -104.5548
17.6243 | -104.554
17.628 | -104.5557
17.6317 | -104.5591
17.6353 | -104.5621
17.639 | -104.5655
17.6427 | -104.5703
17.6464 | -104.5764
17.65 | -104.585
17.6537 | -104.5939
17.6574 | -104.604
17.661 | -104.6142
17.6647 | -104.6237
17.6684 | -104.6323
17.6721 | -104.6381
17.6757 | -104.6451
17.6794 | -104.6539
17.6831 | -104.6627
17.6868 | -104.6724
17.6904 | -104.6807
17.6941 | -104.6869
17.6978 | -104.693
17.7015 | -104.6976
17.7051 | -104.6991
17.7088 | -104.6995
17.7125 | -104.7018
17.7161 | -104.7028
17.7198 | -104.7008
17.7235 | -104.698
17.7272 | -104.6975
17.7308 | -104.6948
17.7345 | -104.689
17.7382 | -104.6798
17.7419 | -104.6703
17.7455 | -104.6587
17.7492 | -104.6472
17.7529 | -104.6371
17.7566 | -104.6256
17.7602 | -104.6157
17.7639 | -104.604
17.7676 | -104.5902
17.7712 | -104.5765
17.7749 | -104.5629
17.7786 | -104.55
17.7823 | -104.5378
17.7859 | -104.5263
17.7896 | -104.5173
17.7933 | -104.5099
17.797 | -104.5021
17.8006 | -104.4961
17.8043 | -104.49
17.808 | -104.4858
17.8117 | -104.4804
17.8153 | -104.473
17.819 | -104.4663
17.8227 | -104.4591
17.8263 | -104.4541
17.83 | -104.4522
17.8337 | -104.4507
17.8374 | -104.4508
17.841 | -104.452
17.8447 | -104.4517
17.8484 | -104.4521
17.8521 | -104.4548
17.8557 | -104.4586
17.8594 | -104.4625
17.8631 | -104.4663
17.8668 | -104.4719
17.8704 | -104.4756
17.8741 | -104.4803
17.8778 | -104.4881
17.8814 | -104.4929
17.8851 | -104.4962
17.8888 | -104.5011
17.8925 | -104.5026
17.8961 | -104.5041
17.8998 | -104.5072
17.9035 | -104.5103
17.9072 | -104.5123
17.9108 | -104.5131
17.9145 | -104.5152
17.9182 | -104.5164
17.9219 | -104.5167
17.9255 | -104.5179
17.9292 | -104.5188
17.9329 | -104.513
17.9365 | -104.5063
17.9402 | -104.5003
17.9439 | -104.4955
17.9476 | -104.4895
17.9512 | -104.4819
17.9549 | -104.4716
17.9586 | -104.4593
17.9623 | -104.441
17.9659 | -104.4185
17.9696 | -104.3933
17.9733 | -104.3693
17.9769 | -104.3467
17.9806 | -104.3208
17.9843 | -104.2954
17.988 | -104.2709
17.9916 | -104.246
17.9953 | -104.219
17.999 | -104.1888
18.0027 | -104.1564
18.0063 | -104.1227
18.01 | -104.0907
18.0137 | -104.0539
18.0174 | -104.0127
18.021 | -103.9764
18.0247 | -103.9402
18.0284 | -103.9049
18.032 | -103.8704
18.0357 | -103.8357
18.0394 | -103.8038
18.0431 | -103.7702
18.0467 | -103.7367
18.0504 | -103.7054
18.0541 | -103.6756
18.0578 | -103.6474
18.0614 | -103.6219
18.0651 | -103.5982
18.0688 | -103.5778
18.0725 | -103.5621
18.0761 | -103.5495
18.0798 | -103.5374
18.0835 | -103.5269
18.0871 | -103.5165
18.0908 | -103.5069
18.0945 | -103.5005
18.0982 | -103.4989
18.1018 | -103.4991
18.1055 | -103.4989
18.1092 | -103.5017
18.1129 | -103.5055
18.1165 | -103.5106
18.1202 | -103.5191
18.1239 | -103.5268
18.1276 | -103.5351
18.1312 | -103.5445
18.1349 | -103.5538
18.1386 | -103.5641
18.1422 | -103.5736
18.1459 | -103.5805
18.1496 | -103.5893
18.1533 | -103.5962
18.1569 | -103.6001
18.1606 | -103.6021
18.1643 | -103.6015
18.168 | -103.5985
18.1716 | -103.596
18.1753 | -103.5922
18.179 | -103.586
18.1827 | -103.5791
18.1863 | -103.5719
18.19 | -103.5658
18.1937 | -103.5583
18.1973 | -103.5498
18.201 | -103.5398
18.2047 | -103.5301
18.2084 | -103.5192
18.212 | -103.5094
18.2157 | -103.5021
18.2194 | -103.4976
18.2231 | -103.4952
18.2267 | -103.4924
18.2304 | -103.4884
18.2341 | -103.4839
18.2378 | -103.4777
18.2414 | -103.4725
18.2451 | -103.47
18.2488 | -103.4678
18.2524 | -103.4674
18.2561 | -103.4684
18.2598 | -103.4713
18.2635 | -103.4756
18.2671 | -103.4771
18.2708 | -103.4813
18.2745 | -103.4882
18.2782 | -103.4959
18.2818 | -103.5052
18.2855 | -103.5148
18.2892 | -103.5235
18.2929 | -103.5308
18.2965 | -103.5366
18.3002 | -103.5432
18.3039 | -103.5505
18.3075 | -103.5588
18.3112 | -103.5662
18.3149 | -103.5733
18.3186 | -103.5801
18.3222 | -103.5854
18.3259 | -103.5919
18.3296 | -103.5981
18.3333 | -103.6018
18.3369 | -103.6047
18.3406 | -103.6061
18.3443 | -103.6051
18.348 | -103.6061
18.3516 | -103.6068
18.3553 | -103.6071
18.359 | -103.604
18.3626 | -103.5999
18.3663 | -103.5961
18.37 | -103.5899
18.3737 | -103.5847
18.3773 | -103.5775
18.381 | -103.5661
18.3847 | -103.5576
18.3884 | -103.5522
18.392 | -103.545
18.3957 | -103.5383
18.3994 | -103.531
18.403 | -103.5235
18.4067 | -103.5186
18.4104 | -103.5154
18.4141 | -103.5124
18.4177 | -103.5086
18.4214 | -103.5053
18.4251 | -103.5055
18.4288 | -103.5083
18.4324 | -103.5086
18.4361 | -103.5093
18.4398 | -103.5126
18.4435 | -103.5171
18.4471 | -103.5191
18.4508 | -103.5234
18.4545 | -103.5315
18.4581 | -103.5391
18.4618 | -103.5491
18.4655 | -103.5606
18.4692 | -103.5727
18.4728 | -103.5861
18.4765 | -103.5993
18.4802 | -103.6124
18.4839 | -103.6269
18.4875 | -103.6428
18.4912 | -103.6599
18.4949 | -103.6767
18.4986 | -103.6961
18.5022 | -103.7148
18.5059 | -103.7308
18.5096 | -103.7486
18.5132 | -103.7684
18.5169 | -103.7887
18.5206 | -103.808
18.5243 | -103.8278
18.5279 | -103.8499
18.5316 | -103.8706
18.5353 | -103.89
18.539 | -103.907
18.5426 | -103.9239
18.5463 | -103.9346
18.55 | -103.9494
18.5537 | -103.9619
18.5573 | -103.9754
18.561 | -103.9888
18.5647 | -104.0012
18.5683 | -104.0122
18.572 | -104.0228
18.5757 | -104.0313
18.5794 | -104.035
18.583 | -104.0393
18.5867 | -104.0433
18.5904 | -104.0455
18.5941 | -104.0472
18.5977 | -104.0494
18.6014 | -104.0526
18.6051 | -104.057
18.6088 | -104.0597
18.6124 | -104.0591
18.6161 | -104.0595
18.6198 | -104.0567
18.6234 | -104.0523
18.6271 | -104.0482
18.6308 | -104.0452
18.6345 | -104.0408
18.6381 | -104.033
18.6418 | -104.0261
18.6455 | -104.0199
18.6492 | -104.011
18.6528 | -103.9995
18.6565 | -103.9874
18.6602 | -103.9752
18.6639 | -103.9641
18.6675 | -103.9524
18.6712 | -103.9387
18.6749 | -103.927
18.6785 | -103.9124
18.6822 | -103.8973
18.6859 | -103.8835
18.6896 | -103.8679
18.6932 | -103.8525
18.6969 | -103.8393
18.7006 | -103.8273
18.7043 | -103.8148
18.7079 | -103.8023
18.7116 | -103.7896
18.7153 | -103.7756
18.7189 | -103.7598
18.7226 | -103.7443
18.7263 | -103.729
18.73 | -103.7143
18.7336 | -103.7014
18.7373 | -103.6877
18.741 | -103.673
18.7447 | -103.6595
18.7483 | -103.6438
18.752 | -103.6309
18.7557 | -103.6166
18.7594 | -103.6037
18.763 | -103.5904
18.7667 | -103.5788
18.7704 | -103.5691
18.774 | -103.5591
18.7777 | -103.5466
18.7814 | -103.5363
18.7851 | -103.5266
18.7887 | -103.5151
18.7924 | -103.5059
18.7961 | -103.4995
18.7998 | -103.4927
18.8034 | -103.4883
18.8071 | -103.4861
18.8108 | -103.4839
18.8145 | -103.4809
18.8181 | -103.4807
18.8218 | -103.4824
18.8255 | -103.4829
18.8291 | -103.4862
18.8328 | -103.4881
18.8365 | -103.4903
18.8402 | -103.4922
18.8438 | -103.4955
18.8475 | -103.5041
18.8512 | -103.5103
18.8549 | -103.5155
18.8585 | -103.519
18.8622 | -103.5231
18.8659 | -103.5288
18.8696 | -103.5311
18.8732 | -103.5357
18.8769 | -103.5392
18.8806 | -103.5418
18.8842 | -103.5451
18.8879 | -103.5503
18.8916 | -103.5534
18.8953 | -103.5557
18.8989 | -103.559
18.9026 | -103.5626
18.9063 | -103.5667
18.91 | -103.57
18.9136 | -103.571
18.9173 | -103.5716
18.921 | -103.5698
18.9247 | -103.5661
18.9283 | -103.5629
18.932 | -103.5593
18.9357 | -103.5557
18.9393 | -103.5493
18.943 | -103.5389
18.9467 | -103.5303
18.9504 | -103.5208
18.954 | -103.5102
18.9577 | -103.4992
18.9614 | -103.488
18.9651 | -103.4747
18.9687 | -103.4612
18.9724 | -103.447
18.9761 | -103.4317
18.9798 | -103.4191
18.9834 | -103.4055
18.9871 | -103.3899
18.9908 | -103.375
18.9944 | -103.3609
18.9981 | -103.3463
19.0018 | -103.3316
19.0055 | -103.317
19.0091 | -103.3008
19.0128 | -103.2854
19.0165 | -103.2718
19.0202 | -103.2554
19.0238 | -103.2427
19.0275 | -103.2277
19.0312 | -103.2141
19.0348 | -103.1993
19.0385 | -103.1847
19.0422 | -103.1703
19.0459 | -103.1579
19.0495 | -103.1463
19.0532 | -103.1374
19.0569 | -103.1289
19.0606 | -103.1207
19.0642 | -103.1129
19.0679 | -103.1046
19.0716 | -103.0961
19.0753 | -103.0952
19.0789 | -103.0868
19.0826 | -103.0783
19.0863 | -103.0749
19.0899 | -103.0713
19.0936 | -103.066
19.0973 | -103.0605
19.101 | -103.0551
19.1046 | -103.0488
19.1083 | -103.0441
19.112 | -103.0413
19.1157 | -103.0392
19.1193 | -103.0351
19.123 | -103.0324
19.1267 | -103.0299
19.1304 | -103.0296
19.134 | -103.0279
19.1377 | -103.0213
19.1414 | -103.0143
19.145 | -103.0052
19.1487 | -102.999
19.1524 | -102.9925
19.1561 | -102.9848
19.1597 | -102.9738
19.1634 | -102.9636
19.1671 | -102.9537
19.1708 | -102.9427
19.1744 | -102.9315
19.1781 | -102.9208
19.1818 | -102.9079
19.1855 | -102.8938
19.1891 | -102.8804
19.1928 | -102.869
19.1965 | -102.8582
19.2001 | -102.8444
19.2038 | -102.8302
19.2075 | -102.8179
19.2112 | -102.8065
19.2148 | -102.795
19.2185 | -102.785
19.2222 | -102.7746
19.2259 | -102.7641
19.2295 | -102.755
19.2332 | -102.7451
19.2369 | -102.735
19.2406 | -102.726
19.2442 | -102.7187
19.2479 | -102.7122
19.2516 | -102.7047
19.2552 | -102.6993
19.2589 | -102.6952
19.2626 | -102.6912
19.2663 | -102.6869
19.2699 | -102.6839
19.2736 | -102.6817
19.2773 | -102.6801
19.281 | -102.6796
19.2846 | -102.678
19.2883 | -102.6803
19.292 | -102.6844
19.2957 | -102.6867
19.2993 | -102.6894
19.303 | -102.6925
19.3067 | -102.6964
19.3103 | -102.6984
19.314 | -102.7014
19.3177 | -102.7053
19.3214 | -102.7064
19.325 | -102.7079
19.3287 | -102.7127
19.3324 | -102.7151
19.3361 | -102.7161
19.3397 | -102.7207
19.3434 | -102.7246
19.3471 | -102.7281
19.3508 | -102.7312
19.3544 | -102.7367
19.3581 | -102.7444
19.3618 | -102.7477
19.3654 | -102.7488
19.3691 | -102.7505
19.3728 | -102.7522
19.3765 | -102.7542
19.3801 | -102.7565
19.3838 | -102.7587
19.3875 | -102.7591
19.3912 | -102.7561
19.3948 | -102.7561
19.3985 | -102.7556
19.4022 | -102.7549
19.4059 | -102.7525
19.4095 | -102.7476
19.4132 | -102.7468
19.4169 | -102.7493
19.4205 | -102.7491
19.4242 | -102.7476
19.4279 | -102.7461
19.4316 | -102.7419
19.4352 | -102.7408
19.4389 | -102.7377
19.4426 | -102.7316
19.4463 | -102.7276
19.4499 | -102.7215
19.4536 | -102.7177
19.4573 | -102.7162
19.4609 | -102.7151
19.4646 | -102.7128
19.4683 | -102.7103
19.472 | -102.7077
19.4756 | -102.7062
19.4793 | -102.7018
19.483 | -102.6975
19.4867 | -102.6939
19.4903 | -102.6903
19.494 | -102.6865
19.4977 | -102.6829
19.5014 | -102.6769
19.505 | -102.6711
19.5087 | -102.6652
19.5124 | -102.662
19.516 | -102.6588
19.5197 | -102.6561
19.5234 | -102.6535
19.5271 | -102.6503
19.5307 | -102.6463
19.5344 | -102.6422
19.5381 | -102.6394
19.5418 | -102.6358
19.5454 | -102.6309
19.5491 | -102.6267
19.5528 | -102.6229
19.5565 | -102.6207
19.5601 | -102.6195
19.5638 | -102.6154
19.5675 | -102.6115
19.5711 | -102.6092
19.5748 | -102.6082
19.5785 | -102.6077
19.5822 | -102.6068
19.5858 | -102.6077
19.5895 | -102.607
19.5932 | -102.606
19.5969 | -102.6072
19.6005 | -102.6077
19.6042 | -102.6116
19.6079 | -102.6117
19.6116 | -102.6122
19.6152 | -102.6149
19.6189 | -102.6181
19.6226 | -102.6207
19.6262 | -102.624
19.6299 | -102.6273
19.6336 | -102.6292
19.6373 | -102.6332
19.6409 | -102.6394
19.6446 | -102.6465
19.6483 | -102.654
19.652 | -102.6639
19.6556 | -102.6735
19.6593 | -102.6818
19.663 | -102.6901
19.6667 | -102.6987
19.6703 | -102.709
19.674 | -102.7165
19.6777 | -102.7247
19.6813 | -102.7351
19.685 | -102.7465
19.6887 | -102.7578
19.6924 | -102.7694
19.696 | -102.7788
19.6997 | -102.7882
19.7034 | -102.797
19.7071 | -102.8056
19.7107 | -102.8157
19.7144 | -102.8256
19.7181 | -102.8374
19.7218 | -102.8487
19.7254 | -102.8555
19.7291 | -102.8643
19.7328 | -102.8715
19.7364 | -102.8783
19.7401 | -102.8837
19.7438 | -102.8881
19.7475 | -102.8924
19.7511 | -102.8957
19.7548 | -102.8981
19.7585 | -102.9042
19.7622 | -102.9081
19.7658 | -102.9112
19.7695 | -102.913
19.7732 | -102.9138
19.7768 | -102.9156
19.7805 | -102.9185
19.7842 | -102.9187
19.7879 | -102.919
19.7915 | -102.9175
19.7952 | -102.9157
19.7989 | -102.9143
19.8026 | -102.9111
19.8062 | -102.9089
19.8099 | -102.904
19.8136 | -102.8969
19.8173 | -102.8905
19.8209 | -102.8838
19.8246 | -102.8751
19.8283 | -102.8655
19.8319 | -102.8539
19.8356 | -102.8419
19.8393 | -102.8291
19.843 | -102.8161
19.8466 | -102.8034
19.8503 | -102.7894
19.854 | -102.7732
19.8577 | -102.7581
19.8613 | -102.7433
19.865 | -102.7273
19.8687 | -102.7132
19.8724 | -102.6984
19.876 | -102.6832
19.8797 | -102.6669
19.8834 | -102.6485
19.887 | -102.6295
19.8907 | -102.6134
19.8944 | -102.6006
19.8981 | -102.5864
19.9017 | -102.5711
19.9054 | -102.5584
19.9091 | -102.5469
19.9128 | -102.533
19.9164 | -102.5209
19.9201 | -102.5097
19.9238 | -102.4986
19.9275 | -102.4886
19.9311 | -102.4808
19.9348 | -102.4713
19.9385 | -102.4625
19.9421 | -102.4578
19.9458 | -102.4525
19.9495 | -102.4468
19.9532 | -102.4427
19.9568 | -102.4397
19.9605 | -102.4355
19.9642 | -102.4317
19.9679 | -102.4282
19.9715 | -102.4265
19.9752 | -102.4226
19.9789 | -102.4201
19.9826 | -102.4205
19.9862 | -102.4204
19.9899 | -102.4198
19.9936 | -102.4176
19.9972 | -102.4147
20.0009 | -102.4133
20.0046 | -102.4135
20.0083 | -102.4156
20.0119 | -102.4183
20.0156 | -102.4198
20.0193 | -102.4223
20.023 | -102.4277
20.0266 | -102.4347
20.0303 | -102.4394
20.034 | -102.4456
20.0377 | -102.4537
20.0413 | -102.4601
20.045 | -102.4673
20.0487 | -102.4751
20.0523 | -102.4851
20.056 | -102.4945
20.0597 | -102.5039
20.0634 | -102.5146
20.067 | -102.5276
20.0707 | -102.5403
20.0744 | -102.5537
20.0781 | -102.5666
20.0817 | -102.5791
20.0854 | -102.5944
20.0891 | -102.6096
20.0927 | -102.6269
20.0964 | -102.6467
20.1001 | -102.6649
20.1038 | -102.682
20.1074 | -102.7
20.1111 | -102.7173
20.1148 | -102.7348
20.1185 | -102.7519
20.1221 | -102.7673
20.1258 | -102.7818
20.1295 | -102.7967
20.1332 | -102.8128
20.1368 | -102.8251
20.1405 | -102.8416
20.1442 | -102.8557
20.1478 | -102.8688
20.1515 | -102.883
20.1552 | -102.8944
20.1589 | -102.9049
20.1625 | -102.9144
20.1662 | -102.9232
20.1699 | -102.9345
20.1736 | -102.9456
20.1772 | -102.951
20.1809 | -102.9532
20.1846 | -102.9561
20.1883 | -102.9562
20.1919 | -102.9558
20.1956 | -102.9536
20.1993 | -102.9538
20.2029 | -102.9525
20.2066 | -102.9478
20.2103 | -102.9428
20.214 | -102.9367
20.2176 | -102.9279
20.2213 | -102.9196
20.225 | -102.9089
20.2287 | -102.8957
20.2323 | -102.8831
20.236 | -102.8713
20.2397 | -102.8586
20.2434 | -102.846
20.247 | -102.8319
20.2507 | -102.8161
20.2544 | -102.7983
20.258 | -102.7792
20.2617 | -102.7619
20.2654 | -102.7444
20.2691 | -102.7252
20.2727 | -102.7056
20.2764 | -102.6846
20.2801 | -102.6677
20.2838 | -102.6493
20.2874 | -102.6285
20.2911 | -102.6064
20.2948 | -102.5851
20.2985 | -102.5627
20.3021 | -102.5406
20.3058 | -102.5182
20.3095 | -102.4964
20.3131 | -102.475
20.3168 | -102.4564
20.3205 | -102.4374
20.3242 | -102.4149
20.3278 | -102.3924
20.3315 | -102.3716
20.3352 | -102.3508
20.3389 | -102.3296
20.3425 | -102.3088
20.3462 | -102.2869
20.3499 | -102.2627
20.3536 | -102.2388
20.3572 | -102.2171
20.3609 | -102.1953
20.3646 | -102.1738
20.3682 | -102.1503
20.3719 | -102.1257
20.3756 | -102.1025
20.3793 | -102.0805
20.3829 | -102.0585
20.3866 | -102.0348
20.3903 | -102.0139
20.394 | -101.9928
20.3976 | -101.973
20.4013 | -101.9533
20.405 | -101.9345
20.4087 | -101.9169
20.4123 | -101.9012
20.416 | -101.8844
20.4197 | -101.867
20.4233 | -101.8475
20.427 | -101.8316
20.4307 | -101.8144
20.4344 | -101.7981
20.438 | -101.7872
20.4417 | -101.7772
20.4454 | -101.7666
20.4491 | -101.7556
20.4527 | -101.747
20.4564 | -101.7406
20.4601 | -101.7334
20.4637 | -101.7268
20.4674 | -101.7226
20.4711 | -101.7194
20.4748 | -101.7166
20.4784 | -101.7147
20.4821 | -101.7104
20.4858 | -101.7057
20.4895 | -101.7018
20.4931 | -101.699
20.4968 | -101.6964
20.5005 | -101.6939
20.5042 | -101.6922
20.5078 | -101.6908
20.5115 | -101.6896
20.5152 | -101.6894
20.5188 | -101.6888
20.5225 | -101.6847
20.5262 | -101.6816
20.5299 | -101.6782
20.5335 | -101.6751
20.5372 | -101.674
20.5409 | -101.6695
20.5446 | -101.665
20.5482 | -101.6612
20.5519 | -101.6532
20.5556 | -101.6433
20.5593 | -101.6332
20.5629 | -101.6232
20.5666 | -101.614
20.5703 | -101.6028
20.5739 | -101.59
20.5776 | -101.5768
20.5813 | -101.5652
20.585 | -101.5498
20.5886 | -101.5329
20.5923 | -101.518
20.596 | -101.5011
20.5997 | -101.4802
20.6033 | -101.4559
20.607 | -101.432
20.6107 | -101.41
20.6144 | -101.3896
20.618 | -101.3681
20.6217 | -101.3458
20.6254 | -101.3185
20.629 | -101.2845
20.6327 | -101.2539
20.6364 | -101.2239
20.6401 | -101.1928
20.6437 | -101.1624
20.6474 | -101.1268
20.6511 | -101.0898
20.6548 | -101.055
20.6584 | -101.0227
20.6621 | -100.9912
20.6658 | -100.954
20.6695 | -100.9201
20.6731 | -100.8846
20.6768 | -100.8505
20.6805 | -100.8184
20.6841 | -100.7855
20.6878 | -100.7521
20.6915 | -100.7194
20.6952 | -100.6888
20.6988 | -100.6573
20.7025 | -100.6248
20.7062 | -100.5944
20.7099 | -100.5675
20.7135 | -100.5414
20.7172 | -100.5178
20.7209 | -100.4919
20.7246 | -100.4678
20.7282 | -100.442
20.7319 | -100.4161
20.7356 | -100.3923
20.7392 | -100.3702
20.7429 | -100.3494
20.7466 | -100.3274
20.7503 | -100.3072
20.7539 | -100.2884
20.7576 | -100.2706
20.7613 | -100.2535
20.765 | -100.2367
20.7686 | -100.2203
20.7723 | -100.2073
20.776 | -100.1962
20.7797 | -100.1867
20.7833 | -100.1769
20.787 | -100.1671
20.7907 | -100.1588
20.7943 | -100.1503
20.798 | -100.1407
20.8017 | -100.132
20.8054 | -100.1255
20.809 | -100.1163
20.8127 | -100.108
20.8164 | -100.1045
20.8201 | -100.1016
20.8237 | -100.1
20.8274 | -100.0979
20.8311 | -100.0956
20.8347 | -100.0947
20.8384 | -100.0961
20.8421 | -100.0967
20.8458 | -100.098
20.8494 | -100.1001
20.8531 | -100.1002
20.8568 | -100.1002
20.8605 | -100.1029
20.8641 | -100.1067
20.8678 | -100.1119
20.8715 | -100.1167
20.8752 | -100.119
20.8788 | -100.1238
20.8825 | -100.133
20.8862 | -100.1419
20.8898 | -100.1523
20.8935 | -100.1654
20.8972 | -100.181
20.9009 | -100.1948
20.9045 | -100.2076
20.9082 | -100.2235
20.9119 | -100.241
20.9156 | -100.258
20.9192 | -100.2789
20.9229 | -100.3049
20.9266 | -100.3278
20.9303 | -100.3511
20.9339 | -100.3761
20.9376 | -100.4013
20.9413 | -100.4303
20.9449 | -100.4586
20.9486 | -100.4845
20.9523 | -100.511
20.956 | -100.5378
20.9596 | -100.5649
20.9633 | -100.5947
20.967 | -100.6234
20.9707 | -100.6527
20.9743 | -100.6807
20.978 | -100.7085
20.9817 | -100.7371
20.9854 | -100.7685
20.989 | -100.7974
20.9927 | -100.8234
20.9964 | -100.85
21 | -100.8781
21.0037 | -100.9036
21.0074 | -100.9263
21.0111 | -100.9498
21.0147 | -100.9739
21.0184 | -100.9978
21.0221 | -101.0216
21.0258 | -101.0469
21.0294 | -101.0674
21.0331 | -101.0883
21.0368 | -101.1112
21.0405 | -101.1326
21.0441 | -101.1533
21.0478 | -101.1725
21.0515 | -101.1917
21.0551 | -101.2102
21.0588 | -101.2265
21.0625 | -101.242
21.0662 | -101.2577
21.0698 | -101.2694
21.0735 | -101.2795
21.0772 | -101.2885
21.0809 | -101.2967
21.0845 | -101.3049
21.0882 | -101.3114
21.0919 | -101.3181
21.0956 | -101.3269
21.0992 | -101.3337
21.1029 | -101.3361
21.1066 | -101.3416
21.1102 | -101.3468
21.1139 | -101.3495
21.1176 | -101.3503
21.1213 | -101.3511
21.1249 | -101.3521
21.1286 | -101.3549
21.1323 | -101.359
21.136 | -101.3594
21.1396 | -101.3622
21.1433 | -101.3656
21.147 | -101.3699
21.1506 | -101.3778
21.1543 | -101.3852
21.158 | -101.3901
21.1617 | -101.3948
21.1653 | -101.4007
21.169 | -101.409
21.1727 | -101.4206
21.1764 | -101.4313
21.18 | -101.4382
21.1837 | -101.4458
21.1874 | -101.4558
21.1911 | -101.4646
21.1947 | -101.4761
21.1984 | -101.4849
21.2021 | -101.4959
21.2057 | -101.5098
21.2094 | -101.5208
21.2131 | -101.5333
21.2168 | -101.5477
21.2204 | -101.5629
21.2241 | -101.5761
21.2278 | -101.5904
21.2315 | -101.6047
21.2351 | -101.6172
21.2388 | -101.6278
21.2425 | -101.6383
21.2462 | -101.6519
21.2498 | -101.6659
21.2535 | -101.6782
21.2572 | -101.6907
21.2608 | -101.7026
21.2645 | -101.7125
21.2682 | -101.7228
21.2719 | -101.7349
21.2755 | -101.7453
21.2792 | -101.7545
21.2829 | -101.7639
21.2866 | -101.7724
21.2902 | -101.7838
21.2939 | -101.7956
21.2976 | -101.8076
21.3013 | -101.8161
21.3049 | -101.8227
21.3086 | -101.8293
21.3123 | -101.8355
21.3159 | -101.8426
21.3196 | -101.8508
21.3233 | -101.8572
21.327 | -101.8628
21.3306 | -101.8693
21.3343 | -101.8721
21.338 | -101.8726
21.3417 | -101.8719
21.3453 | -101.8721
21.349 | -101.8713
21.3527 | -101.8697
21.3564 | -101.8693
21.36 | -101.8703
21.3637 | -101.8703
21.3674 | -101.8682
21.371 | -101.8673
21.3747 | -101.8649
21.3784 | -101.8622
21.3821 | -101.8589
21.3857 | -101.853
21.3894 | -101.8472
21.3931 | -101.8403
21.3968 | -101.8339
21.4004 | -101.8278
21.4041 | -101.8223
21.4078 | -101.8141
21.4115 | -101.8063
21.4151 | -101.7987
21.4188 | -101.7906
21.4225 | -101.7818
21.4261 | -101.7747
21.4298 | -101.7654
21.4335 | -101.7557
21.4372 | -101.749
21.4408 | -101.7434
21.4445 | -101.7381
21.4482 | -101.7313
21.4519 | -101.7251
21.4555 | -101.7192
21.4592 | -101.712
21.4629 | -101.7051
21.4666 | -101.698
21.4702 | -101.692
21.4739 | -101.6872
21.4776 | -101.6803
21.4812 | -101.6754
21.4849 | -101.6713
21.4886 | -101.6667
21.4923 | -101.6648
21.4959 | -101.6651
21.4996 | -101.6644
21.5033 | -101.6638
21.507 | -101.6642
21.5106 | -101.6636
21.5143 | -101.664
21.518 | -101.6655
21.5216 | -101.6674
21.5253 | -101.6699
21.529 | -101.6743
21.5327 | -101.6788
21.5363 | -101.6861
21.54 | -101.6964
21.5437 | -101.7065
21.5474 | -101.7149
21.551 | -101.722
21.5547 | -101.7307
21.5584 | -101.7402
21.5621 | -101.7504
21.5657 | -101.7652
21.5694 | -101.7811
21.5731 | -101.7978
21.5767 | -101.814
21.5804 | -101.8323
21.5841 | -101.851
21.5878 | -101.8689
21.5914 | -101.8864
21.5951 | -101.9023
21.5988 | -101.9196
21.6025 | -101.9373
21.6061 | -101.9581
21.6098 | -101.9802
21.6135 | -102.002
21.6172 | -102.0214
21.6208 | -102.0382
21.6245 | -102.0523
21.6282 | -102.0687
21.6318 | -102.084
21.6355 | -102.0963
21.6392 | -102.1123
21.6429 | -102.1275
21.6465 | -102.1434
21.6502 | -102.159
21.6539 | -102.174
21.6576 | -102.1868
21.6612 | -102.1972
21.6649 | -102.2072
21.6686 | -102.2185
21.6723 | -102.2292
21.6759 | -102.2376
21.6796 | -102.2468
21.6833 | -102.2536
21.6869 | -102.2602
21.6906 | -102.2696
21.6943 | -102.2776
21.698 | -102.2857
21.7016 | -102.2914
21.7053 | -102.2979
21.709 | -102.3036
21.7127 | -102.3099
21.7163 | -102.3155
21.72 | -102.3201
21.7237 | -102.3243
21.7274 | -102.3291
21.731 | -102.3354
21.7347 | -102.3414
21.7384 | -102.3501
21.742 | -102.3596
21.7457 | -102.3655
21.7494 | -102.371
21.7531 | -102.3782
21.7567 | -102.3868
21.7604 | -102.3948
21.7641 | -102.4024
21.7678 | -102.4144
21.7714 | -102.4277
21.7751 | -102.4431
21.7788 | -102.4571
21.7825 | -102.4699
21.7861 | -102.4848
21.7898 | -102.4993
21.7935 | -102.5157
21.7971 | -102.532
21.8008 | -102.5466
21.8045 | -102.5576
21.8082 | -102.5718
21.8118 | -102.5887
21.8155 | -102.6036
21.8192 | -102.6179
21.8229 | -102.6333
21.8265 | -102.6466
21.8302 | -102.6602
21.8339 | -102.675
21.8376 | -102.6891
21.8412 | -102.7016
21.8449 | -102.7121
21.8486 | -102.7209
21.8522 | -102.7306
21.8559 | -102.7424
21.8596 | -102.7514
21.8633 | -102.7572
21.8669 | -102.7618
21.8706 | -102.7654
21.8743 | -102.7653
21.878 | -102.7664
21.8816 | -102.7673
21.8853 | -102.7664
21.889 | -102.7647
21.8926 | -102.7643
21.8963 | -102.7637
21.9 | -102.7633
21.9037 | -102.764
21.9073 | -102.7648
21.911 | -102.7633
21.9147 | -102.7616
21.9184 | -102.7599
21.922 | -102.7579
21.9257 | -102.7571
21.9294 | -102.7546
21.9331 | -102.7518
21.9367 | -102.7469
21.9404 | -102.7428
21.9441 | -102.7409
21.9477 | -102.7364
21.9514 | -102.7322
21.9551 | -102.7277
21.9588 | -102.7217
21.9624 | -102.7146
21.9661 | -102.7079
21.9698 | -102.6997
21.9735 | -102.6938
21.9771 | -102.6888
21.9808 | -102.6848
21.9845 | -102.6811
21.9882 | -102.6768
21.9918 | -102.6739
21.9955 | -102.6689
21.9992 | -102.6638
22.0028 | -102.6612
22.0065 | -102.6579
22.0102 | -102.6555
22.0139 | -102.6519
22.0175 | -102.6499
22.0212 | -102.6502
22.0249 | -102.6482
22.0286 | -102.6465
22.0322 | -102.6459
22.0359 | -102.6481
22.0396 | -102.6493
22.0433 | -102.6498
22.0469 | -102.6499
22.0506 | -102.6511
22.0543 | -102.6517
22.0579 | -102.6555
22.0616 | -102.6583
22.0653 | -102.6611
22.069 | -102.6641
22.0726 | -102.6674
22.0763 | -102.6726
22.08 | -102.676
22.0837 | -102.6782
22.0873 | -102.68
22.091 | -102.6853
22.0947 | -102.6896
22.0984 | -102.6952
22.102 | -102.6993
22.1057 | -102.7022
22.1094 | -102.7074
22.113 | -102.7138
22.1167 | -102.723
22.1204 | -102.7297
22.1241 | -102.733
22.1277 | -102.7382
22.1314 | -102.7455
22.1351 | -102.7533
22.1388 | -102.7595
22.1424 | -102.765
22.1461 | -102.7712
22.1498 | -102.7762
22.1535 | -102.7816
22.1571 | -102.787
22.1608 | -102.7932
22.1645 | -102.7997
22.1681 | -102.8046
22.1718 | -102.809
22.1755 | -102.8119
22.1792 | -102.8156
22.1828 | -102.8183
22.1865 | -102.8211
22.1902 | -102.8224
22.1939 | -102.8226
22.1975 | -102.824
22.2012 | -102.826
22.2049 | -102.8267
22.2085 | -102.8282
22.2122 | -102.8305
22.2159 | -102.8325
22.2196 | -102.8333
22.2232 | -102.8351
22.2269 | -102.837
22.2306 | -102.8392
22.2343 | -102.8429
22.2379 | -102.845
22.2416 | -102.8498
22.2453 | -102.8555
22.249 | -102.8612
22.2526 | -102.866
22.2563 | -102.876
22.26 | -102.8832
22.2636 | -102.8884
22.2673 | -102.8943
22.271 | -102.9016
22.2747 | -102.9102
22.2783 | -102.9195
22.282 | -102.9302
22.2857 | -102.9421
22.2894 | -102.9513
22.293 | -102.9601
22.2967 | -102.9678
22.3004 | -102.9759
22.3041 | -102.9853
22.3077 | -102.9923
22.3114 | -102.9996
22.3151 | -103.0072
22.3187 | -103.015
22.3224 | -103.0222
22.3261 | -103.0279
22.3298 | -103.0348
22.3334 | -103.043
22.3371 | -103.0502
22.3408 | -103.0584
22.3445 | -103.0646
22.3481 | -103.0703
22.3518 | -103.0737
22.3555 | -103.0781
22.3592 | -103.0822
22.3628 | -103.0875
22.3665 | -103.0909
22.3702 | -103.0946
22.3738 | -103.0961
22.3775 | -103.0966
22.3812 | -103.0988
22.3849 | -103.1008
22.3885 | -103.1017
22.3922 | -103.1014
22.3959 | -103.0996
22.3996 | -103.1002
22.4032 | -103.1018
22.4069 | -103.1001
22.4106 | -103.1024
22.4143 | -103.1057
22.4179 | -103.1081
22.4216 | -103.1091
22.4253 | -103.11
22.4289 | -103.1128
22.4326 | -103.117
22.4363 | -103.1208
22.44 | -103.1256
22.4436 | -103.1292
22.4473 | -103.1323
22.451 | -103.1349
22.4547 | -103.1392
22.4583 | -103.1458
22.462 | -103.1498
22.4657 | -103.1528
22.4694 | -103.1564
22.473 | -103.1607
22.4767 | -103.1638
22.4804 | -103.1658
22.484 | -103.1716
22.4877 | -103.1799
22.4914 | -103.1896
22.4951 | -103.1981
22.4987 | -103.2031
22.5024 | -103.2047
22.5061 | -103.2065
22.5098 | -103.2096
22.5134 | -103.2111
22.5171 | -103.2116
22.5208 | -103.2112
22.5245 | -103.2092
22.5281 | -103.2053
22.5318 | -103.1988
22.5355 | -103.1922
22.5391 | -103.1856
22.5428 | -103.1769
22.5465 | -103.166
22.5502 | -103.1537
22.5538 | -103.1371
22.5575 | -103.1144
22.5612 | -103.0925
22.5649 | -103.0655
22.5685 | -103.0362
22.5722 | -103.0048
22.5759 | -102.9715
22.5795 | -102.9398
22.5832 | -102.9069
22.5869 | -102.8699
22.5906 | -102.8345
22.5942 | -102.797
22.5979 | -102.7568
22.6016 | -102.716
22.6053 | -102.6746
22.6089 | -102.6327
22.6126 | -102.5918
22.6163 | -102.5497
22.62 | -102.5066
22.6236 | -102.4628
22.6273 | -102.4204
22.631 | -102.3775
22.6346 | -102.3324
22.6383 | -102.2875
22.642 | -102.2448
22.6457 | -102.2031
22.6493 | -102.1614
22.653 | -102.1197
22.6567 | -102.0792
22.6604 | -102.0397
22.664 | -102.0035
22.6677 | -101.9695
22.6714 | -101.9367
22.6751 | -101.9041
22.6787 | -101.8725
22.6824 | -101.8434
22.6861 | -101.8167
22.6897 | -101.7921
22.6934 | -101.7708
22.6971 | -101.752
22.7008 | -101.7332
22.7044 | -101.7157
22.7081 | -101.7018
22.7118 | -101.6898
22.7155 | -101.6822
22.7191 | -101.674
22.7228 | -101.6658
22.7265 | -101.6595
22.7302 | -101.6549
22.7338 | -101.6532
22.7375 | -101.6554
22.7412 | -101.6593
22.7448 | -101.6639
22.7485 | -101.6704
22.7522 | -101.6785
22.7559 | -101.689
22.7595 | -101.7005
22.7632 | -101.7144
22.7669 | -101.7267
22.7706 | -101.7403
22.7742 | -101.7563
22.7779 | -101.773
22.7816 | -101.7916
22.7853 | -101.8082
22.7889 | -101.8241
22.7926 | -101.8423
22.7963 | -101.8579
22.7999 | -101.874
22.8036 | -101.8911
22.8073 | -101.9063
22.811 | -101.921
22.8146 | -101.9341
22.8183 | -101.9464
22.822 | -101.9594
22.8257 | -101.9721
22.8293 | -101.9833
22.833 | -101.9936
22.8367 | -102.0031
22.8404 | -102.0121
22.844 | -102.0217
22.8477 | -102.031
22.8514 | -102.04
22.855 | -102.0476
22.8587 | -102.053
22.8624 | -102.0575
22.8661 | -102.0619
22.8697 | -102.0647
22.8734 | -102.0656
22.8771 | -102.0699
22.8808 | -102.0763
22.8844 | -102.0802
22.8881 | -102.0852
22.8918 | -102.0883
22.8955 | -102.0874
22.8991 | -102.0859
22.9028 | -102.0802
22.9065 | -102.075
22.9101 | -102.0706
22.9138 | -102.0643
22.9175 | -102.0594
22.9212 | -102.0571
22.9248 | -102.0542
22.9285 | -102.0495
22.9322 | -102.0451
22.9359 | -102.0395
22.9395 | -102.0352
22.9432 | -102.0315
22.9469 | -102.0255
22.9505 | -102.0207
22.9542 | -102.0168
22.9579 | -102.0133
22.9616 | -102.0087
22.9652 | -102.0039
22.9689 | -102.0005
22.9726 | -101.9972
22.9763 | -101.9907
22.9799 | -101.9842
22.9836 | -101.9792
22.9873 | -101.9744
22.991 | -101.9679
22.9946 | -101.9628
22.9983 | -101.958
23.002 | -101.9534
23.0056 | -101.9488
23.0093 | -101.9453
23.013 | -101.9424
23.0167 | -101.9399
23.0203 | -101.9361
23.024 | -101.9332
23.0277 | -101.9307
23.0314 | -101.9293
23.035 | -101.9286
23.0387 | -101.9264
23.0424 | -101.9244
23.0461 | -101.9245
23.0497 | -101.9262
23.0534 | -101.9266
23.0571 | -101.9261
23.0607 | -101.9263
23.0644 | -101.9264
23.0681 | -101.9256
23.0718 | -101.9248
23.0754 | -101.9255
23.0791 | -101.9268
23.0828 | -101.9281
23.0865 | -101.9282
23.0901 | -101.9317
23.0938 | -101.9346
23.0975 | -101.9372
23.1012 | -101.9394
23.1048 | -101.9391
23.1085 | -101.9407
23.1122 | -101.9445
23.1158 | -101.9483
23.1195 | -101.9482
23.1232 | -101.9496
23.1269 | -101.95
23.1305 | -101.9484
23.1342 | -101.946
23.1379 | -101.9446
23.1416 | -101.9426
23.1452 | -101.9389
23.1489 | -101.9346
23.1526 | -101.9324
23.1563 | -101.9285
23.1599 | -101.9222
23.1636 | -101.9173
23.1673 | -101.9122
23.1709 | -101.9052
23.1746 | -101.896
23.1783 | -101.8858
23.182 | -101.8763
23.1856 | -101.8667
23.1893 | -101.8561
23.193 | -101.844
23.1967 | -101.8324
23.2003 | -101.8229
23.204 | -101.8108
23.2077 | -101.7976
23.2113 | -101.7849
23.215 | -101.7724
23.2187 | -101.7629
23.2224 | -101.7527
23.226 | -101.7404
23.2297 | -101.731
23.2334 | -101.7213
23.2371 | -101.7117
23.2407 | -101.7012
23.2444 | -101.692
23.2481 | -101.6812
23.2518 | -101.6724
23.2554 | -101.6668
23.2591 | -101.6621
23.2628 | -101.6552
23.2664 | -101.6501
23.2701 | -101.6472
23.2738 | -101.6455
23.2775 | -101.641
23.2811 | -101.6372
23.2848 | -101.6346
23.2885 | -101.6333
23.2922 | -101.6327
23.2958 | -101.6308
23.2995 | -101.6299
23.3032 | -101.6336
23.3069 | -101.6388
23.3105 | -101.6431
23.3142 | -101.6526
23.3179 | -101.6529
23.3215 | -101.654
23.3252 | -101.6561
23.3289 | -101.6573
23.3326 | -101.66
23.3362 | -101.6633
23.3399 | -101.6661
23.3436 | -101.6699
23.3473 | -101.6736
23.3509 | -101.6786
23.3546 | -101.6839
23.3583 | -101.6902
23.362 | -101.6938
23.3656 | -101.6976
23.3693 | -101.702
23.373 | -101.7083
23.3766 | -101.7136
23.3803 | -101.7181
23.384 | -101.7219
23.3877 | -101.7252
23.3913 | -101.7284
23.395 | -101.7318
23.3987 | -101.7367
23.4024 | -101.7406
23.406 | -101.7424
23.4097 | -101.7431
23.4134 | -101.7467
23.4171 | -101.7493
23.4207 | -101.75
23.4244 | -101.7509
23.4281 | -101.7516
23.4317 | -101.7508
23.4354 | -101.7516
23.4391 | -101.7521
23.4428 | -101.7519
23.4464 | -101.7533
23.4501 | -101.755
23.4538 | -101.7564
23.4575 | -101.7582
23.4611 | -101.7606
23.4648 | -101.7629
23.4685 | -101.7649
23.4722 | -101.7688
23.4758 | -101.7733
23.4795 | -101.7776
23.4832 | -101.7824
23.4868 | -101.7882
23.4905 | -101.795
23.4942 | -101.8008
23.4979 | -101.807
23.5015 | -101.814
23.5052 | -101.822
23.5089 | -101.832
23.5126 | -101.8418
23.5162 | -101.8528
23.5199 | -101.8644
23.5236 | -101.8767
23.5273 | -101.8889
23.5309 | -101.8994
23.5346 | -101.9112
23.5383 | -101.9231
23.5419 | -101.9345
23.5456 | -101.9466
23.5493 | -101.9585
23.553 | -101.9714
23.5566 | -101.9846
23.5603 | -101.9956
23.564 | -102.0072
23.5677 | -102.0194
23.5713 | -102.0324
23.575 | -102.043
23.5787 | -102.0535
23.5824 | -102.0627
23.586 | -102.0727
23.5897 | -102.082
23.5934 | -102.0944
23.597 | -102.1047
23.6007 | -102.1129
23.6044 | -102.1235
23.6081 | -102.1354
23.6117 | -102.1456
23.6154 | -102.155
23.6191 | -102.1626
23.6228 | -102.1687
23.6264 | -102.1777
23.6301 | -102.1852
23.6338 | -102.1907
23.6374 | -102.1945
23.6411 | -102.1981
23.6448 | -102.2011
23.6485 | -102.2038
23.6521 | -102.2096
23.6558 | -102.2157
23.6595 | -102.2186
23.6632 | -102.2212
23.6668 | -102.2196
23.6705 | -102.2154
23.6742 | -102.2136
23.6779 | -102.2113
23.6815 | -102.2072
23.6852 | -102.2029
23.6889 | -102.1986
23.6925 | -102.1911
23.6962 | -102.1847
23.6999 | -102.1796
23.7036 | -102.1738
23.7072 | -102.1665
23.7109 | -102.1576
23.7146 | -102.1494
23.7183 | -102.1423
23.7219 | -102.1322
23.7256 | -102.1209
23.7293 | -102.1098
23.733 | -102.0987
23.7366 | -102.0902
23.7403 | -102.0829
23.744 | -102.0741
23.7476 | -102.0623
23.7513 | -102.0515
23.755 | -102.0397
23.7587 | -102.0252
23.7623 | -102.0124
23.766 | -102.0006
23.7697 | -101.9909
23.7734 | -101.9803
23.777 | -101.9701
23.7807 | -101.9584
23.7844 | -101.9482
23.7881 | -101.9375
23.7917 | -101.9257
23.7954 | -101.9164
23.7991 | -101.9088
23.8027 | -101.8993
23.8064 | -101.8902
23.8101 | -101.8812
23.8138 | -101.8716
23.8174 | -101.8614
23.8211 | -101.8539
23.8248 | -101.8469
23.8285 | -101.8387
23.8321 | -101.8319
23.8358 | -101.8269
23.8395 | -101.8221
23.8432 | -101.8167
23.8468 | -101.8125
23.8505 | -101.808
23.8542 | -101.8016
23.8578 | -101.7953
23.8615 | -101.7892
23.8652 | -101.786
23.8689 | -101.7839
23.8725 | -101.7806
23.8762 | -101.7757
23.8799 | -101.7711
23.8836 | -101.7678
23.8872 | -101.7649
23.8909 | -101.7636
23.8946 | -101.7634
23.8983 | -101.7624
23.9019 | -101.7608
23.9056 | -101.7609
23.9093 | -101.7599
23.9129 | -101.7595
23.9166 | -101.7615
23.9203 | -101.7624
23.924 | -101.7605
23.9276 | -101.7602
23.9313 | -101.7594
23.935 | -101.7591
23.9387 | -101.7603
23.9423 | -101.7611
23.946 | -101.7611
23.9497 | -101.7626
23.9534 | -101.7641
23.957 | -101.7665
23.9607 | -101.7706
23.9644 | -101.773
23.968 | -101.7763
23.9717 | -101.7796
23.9754 | -101.7838
23.9791 | -101.7887
23.9827 | -101.7918
23.9864 | -101.795
23.9901 | -101.796
23.9938 | -101.7971
23.9974 | -101.8007
24.0011 | -101.8057
24.0048 | -101.8096
24.0084 | -101.8148
24.0121 | -101.8182
24.0158 | -101.8207
24.0195 | -101.8249
24.0231 | -101.8301
24.0268 | -101.8327
24.0305 | -101.8377
24.0342 | -101.8416
24.0378 | -101.8433
24.0415 | -101.8453
24.0452 | -101.8472
24.0489 | -101.8475
24.0525 | -101.8482
24.0562 | -101.8492
24.0599 | -101.8502
24.0635 | -101.851
24.0672 | -101.8534
24.0709 | -101.8536
24.0746 | -101.8488
24.0782 | -101.8454
24.0819 | -101.8432
24.0856 | -101.8405
24.0893 | -101.8377
24.0929 | -101.8364
24.0966 | -101.8339
24.1003 | -101.8282
24.104 | -101.8222
24.1076 | -101.8131
24.1113 | -101.8022
24.115 | -101.7894
24.1186 | -101.7775
24.1223 | -101.7663
24.126 | -101.7535
24.1297 | -101.7398
24.1333 | -101.7256
24.137 | -101.7098
24.1407 | -101.696
24.1444 | -101.6813
24.148 | -101.6666
24.1517 | -101.6531
24.1554 | -101.6393
24.1591 | -101.6221
24.1627 | -101.6037
24.1664 | -101.5874
24.1701 | -101.5708
24.1737 | -101.5507
24.1774 | -101.5334
24.1811 | -101.5164
24.1848 | -101.4998
24.1884 | -101.4842
24.1921 | -101.4638
24.1958 | -101.4436
24.1995 | -101.4252
24.2031 | -101.4095
24.2068 | -101.3899
24.2105 | -101.3683
24.2142 | -101.3514
24.2178 | -101.3352
24.2215 | -101.3191
24.2252 | -101.3025
24.2288 | -101.286
24.2325 | -101.2696
24.2362 | -101.2538
24.2399 | -101.2386
24.2435 | -101.2236
24.2472 | -101.2099
24.2509 | -101.1973
24.2546 | -101.185
24.2582 | -101.1731
24.2619 | -101.162
24.2656 | -101.1531
24.2692 | -101.1447
24.2729 | -101.1373
24.2766 | -101.1299
24.2803 | -101.12
24.2839 | -101.1081
24.2876 | -101.0978
24.2913 | -101.0878
24.295 | -101.08
24.2986 | -101.0732
24.3023 | -101.0656
24.306 | -101.0593
24.3097 | -101.0549
24.3133 | -101.0501
24.317 | -101.0456
24.3207 | -101.0425
24.3243 | -101.0422
24.328 | -101.042
24.3317 | -101.0432
24.3354 | -101.043
24.339 | -101.0412
24.3427 | -101.0397
24.3464 | -101.0411
24.3501 | -101.0423
24.3537 | -101.0425
24.3574 | -101.0429
24.3611 | -101.0438
24.3648 | -101.0434
24.3684 | -101.0463
24.3721 | -101.0499
24.3758 | -101.057
24.3794 | -101.0621
24.3831 | -101.0662
24.3868 | -101.0687
24.3905 | -101.0728
24.3941 | -101.079
24.3978 | -101.0853
24.4015 | -101.0922
24.4052 | -101.0986
24.4088 | -101.1057
24.4125 | -101.1109
24.4162 | -101.1166
24.4199 | -101.1248
24.4235 | -101.1328
24.4272 | -101.1422
24.4309 | -101.1522
24.4345 | -101.163
24.4382 | -101.1728
24.4419 | -101.184
24.4456 | -101.1965
24.4492 | -101.2082
24.4529 | -101.2201
24.4566 | -101.2345
24.4603 | -101.2511
24.4639 | -101.2666
24.4676 | -101.2784
24.4713 | -101.2923
24.475 | -101.3081
24.4786 | -101.3225
24.4823 | -101.3377
24.486 | -101.3523
24.4896 | -101.3679
24.4933 | -101.3854
24.497 | -101.4026
24.5007 | -101.4203
24.5043 | -101.4379
24.508 | -101.4561
24.5117 | -101.4738
24.5154 | -101.4916
24.519 | -101.5113
24.5227 | -101.5287
24.5264 | -101.5462
24.5301 | -101.5629
24.5337 | -101.5807
24.5374 | -101.6023
24.5411 | -101.6225
24.5447 | -101.639
24.5484 | -101.6538
24.5521 | -101.6715
24.5558 | -101.6899
24.5594 | -101.7044
24.5631 | -101.7165
24.5668 | -101.7258
24.5705 | -101.7343
24.5741 | -101.7431
24.5778 | -101.75
24.5815 | -101.7518
24.5852 | -101.7498
24.5888 | -101.7068
24.5925 | -101.6131
24.5962 | -101.5168
24.5998 | -101.4612
24.6035 | -101.4176
24.6072 | -101.375
24.6109 | -101.3242
24.6145 | -101.2649
24.6182 | -101.2004
24.6219 | -101.1329
24.6256 | -101.0652
24.6292 | -100.9975
24.6329 | -100.926
24.6366 | -100.8491
24.6403 | -100.766
24.6439 | -100.6777
24.6476 | -100.5869
24.6513 | -100.4932
24.6549 | -100.3956
24.6586 | -100.2957
24.6623 | -100.1928
24.666 | -100.0868
24.6696 | -99.9805
24.6733 | -99.8751
24.677 | -99.7701
24.6807 | -99.6662
24.6843 | -99.5624
24.688 | -99.4567
24.6917 | -99.351
24.6953 | -99.2477
24.699 | -99.1455
24.7027 | -99.047
24.7064 | -98.9521
24.71 | -98.8616
24.7137 | -98.7738
24.7174 | -98.6889
24.7211 | -98.6086
24.7247 | -98.5327
24.7284 | -98.4622
24.7321 | -98.3967
24.7358 | -98.3366
24.7394 | -98.2807
24.7431 | -98.2296
24.7468 | -98.1835
24.7504 | -98.1429
24.7541 | -98.1091
24.7578 | -98.0803
24.7615 | -98.0578
24.7651 | -98.0411
24.7688 | -98.0304
24.7725 | -98.0247
24.7762 | -98.0245
24.7798 | -98.0299
24.7835 | -98.0416
24.7872 | -98.059
24.7909 | -98.0817
24.7945 | -98.109
24.7982 | -98.1418
24.8019 | -98.1791
24.8055 | -98.2208
24.8092 | -98.2676
24.8129 | -98.3189
24.8166 | -98.3756
24.8202 | -98.4362
24.8239 | -98.4993
24.8276 | -98.5647
24.8313 | -98.6336
24.8349 | -98.7055
24.8386 | -98.7803
24.8423 | -98.8564
24.846 | -98.934
24.8496 | -99.013
24.8533 | -99.0949
24.857 | -99.178
24.8606 | -99.2618
24.8643 | -99.3458
24.868 | -99.4299
24.8717 | -99.5124
24.8753 | -99.5945
24.879 | -99.6759
24.8827 | -99.7551
24.8864 | -99.8322
24.89 | -99.9083
24.8937 | -99.9821
24.8974 | -100.0545
24.9011 | -100.1231
24.9047 | -100.1603
24.9084 | -100.2241
24.9121 | -100.2864
24.9157 | -100.3479
24.9194 | -100.407
24.9231 | -100.4609
24.9268 | -100.5109
24.9304 | -100.5605
24.9341 | -100.6099
24.9378 | -100.6588
24.9415 | -100.7307
24.9451 | -100.8262
24.9488 | -100.9087
24.9525 | -100.9461
24.9562 | -100.9646
24.9598 | -100.9784
24.9635 | -100.9949
24.9672 | -101.012
24.9708 | -101.0297
24.9745 | -101.044
24.9782 | -101.0546
24.9819 | -101.0617
24.9855 | -101.068
24.9892 | -101.0738
24.9929 | -101.0788
24.9966 | -101.0857
25.0002 | -101.0927
25.0039 | -101.0981
25.0076 | -101.1038
25.0112 | -101.1109
25.0149 | -101.1178
25.0186 | -101.1258
25.0223 | -101.1334
25.0259 | -101.1377
25.0296 | -101.1413
25.0333 | -101.146
25.037 | -101.1506
25.0406 | -101.1591
25.0443 | -101.1679
25.048 | -101.1731
25.0517 | -101.181
25.0553 | -101.1882
25.059 | -101.1916
25.0627 | -101.1947
25.0663 | -101.2006
25.07 | -101.2077
25.0737 | -101.2152
25.0774 | -101.2213
25.081 | -101.2251
25.0847 | -101.2306
25.0884 | -101.2379
25.0921 | -101.2454
25.0957 | -101.2536
25.0994 | -101.2601
25.1031 | -101.2662
25.1068 | -101.2722
25.1104 | -101.2795
25.1141 | -101.2869
25.1178 | -101.295
25.1214 | -101.3017
25.1251 | -101.3061
25.1288 | -101.3097
25.1325 | -101.3158
25.1361 | -101.3192
25.1398 | -101.3221
25.1435 | -101.3271
25.1472 | -101.3324
25.1508 | -101.335
25.1545 | -101.3381
25.1582 | -101.3413
25.1619 | -101.3409
25.1655 | -101.3403
25.1692 | -101.339
25.1729 | -101.3374
25.1765 | -101.3369
25.1802 | -101.3372
25.1839 | -101.3339
25.1876 | -101.3298
25.1912 | -101.3242
25.1949 | -101.3186
25.1986 | -101.3144
25.2023 | -101.3091
25.2059 | -101.3012
25.2096 | -101.2911
25.2133 | -101.2817
25.217 | -101.2727
25.2206 | -101.2649
25.2243 | -101.258
25.228 | -101.248
25.2316 | -101.2368
25.2353 | -101.2287
25.239 | -101.2205
25.2427 | -101.2088
25.2463 | -101.198
25.25 | -101.1854
25.2537 | -101.175
25.2574 | -101.1636
25.261 | -101.1511
25.2647 | -101.14
25.2684 | -101.1291
25.2721 | -101.1167
25.2757 | -101.1052
25.2794 | -101.0959
25.2831 | -101.0857
25.2867 | -101.0727
25.2904 | -101.0601
25.2941 | -101.0473
25.2978 | -101.0352
25.3014 | -101.024
25.3051 | -101.0114
25.3088 | -100.9993
25.3125 | -100.9873
25.3161 | -100.9747
25.3198 | -100.9595
25.3235 | -100.9442
25.3271 | -100.9289
25.3308 | -100.9147
25.3345 | -100.9005
25.3382 | -100.8888
25.3418 | -100.8763
25.3455 | -100.8648
25.3492 | -100.8544
25.3529 | -100.8408
25.3565 | -100.8285
25.3602 | -100.8164
25.3639 | -100.8053
25.3676 | -100.7964
25.3712 | -100.7882
25.3749 | -100.7794
25.3786 | -100.7678
25.3822 | -100.7584
25.3859 | -100.7501
25.3896 | -100.7426
25.3933 | -100.7361
25.3969 | -100.7284
25.4006 | -100.7215
25.4043 | -100.7164
25.408 | -100.7111
25.4116 | -100.7033
25.4153 | -100.6974
25.419 | -100.6914
25.4227 | -100.6836
25.4263 | -100.6761
25.43 | -100.6685
25.4337 | -100.6639
25.4373 | -100.6581
25.441 | -100.6532
25.4447 | -100.6494
25.4484 | -100.6436
25.452 | -100.6364
25.4557 | -100.6306
25.4594 | -100.6265
25.4631 | -100.622
25.4667 | -100.6172
25.4704 | -100.6122
25.4741 | -100.6059
25.4778 | -100.6014
25.4814 | -100.5975
25.4851 | -100.5919
25.4888 | -100.588
25.4924 | -100.5843
25.4961 | -100.5807
25.4998 | -100.5743
25.5035 | -100.5682
25.5071 | -100.5608
25.5108 | -100.5542
25.5145 | -100.5522
25.5182 | -100.5492
25.5218 | -100.5459
25.5255 | -100.54
25.5292 | -100.5345
25.5329 | -100.5293
25.5365 | -100.5245
25.5402 | -100.5206
25.5439 | -100.5176
25.5475 | -100.5146
25.5512 | -100.5107
25.5549 | -100.5051
25.5586 | -100.4987
25.5622 | -100.4924
25.5659 | -100.4864
25.5696 | -100.4827
25.5733 | -100.4772
25.5769 | -100.4712
25.5806 | -100.4648
25.5843 | -100.4585
25.588 | -100.4509
25.5916 | -100.4446
25.5953 | -100.4366
25.599 | -100.4285
25.6026 | -100.4241
25.6063 | -100.4194
25.61 | -100.4137
25.6137 | -100.4079
25.6173 | -100.4021
25.621 | -100.397
25.6247 | -100.3903
25.6284 | -100.3829
25.632 | -100.374
25.6357 | -100.3654
25.6394 | -100.3585
25.6431 | -100.353
25.6467 | -100.3489
25.6504 | -100.3457
25.6541 | -100.3418
25.6577 | -100.338
25.6614 | -100.3359
25.6651 | -100.3338
25.6688 | -100.3283
25.6724 | -100.3254
25.6761 | -100.3231
25.6798 | -100.3217
25.6835 | -100.3227
25.6871 | -100.3219
25.6908 | -100.3189
25.6945 | -100.3187
25.6981 | -100.3177
25.7018 | -100.3164
25.7055 | -100.3173
25.7092 | -100.3183
25.7128 | -100.3214
25.7165 | -100.3248
25.7202 | -100.3277
25.7239 | -100.3311
25.7275 | -100.3329
25.7312 | -100.3358
25.7349 | -100.3384
25.7386 | -100.3443
25.7422 | -100.3493
25.7459 | -100.3525
25.7496 | -100.3548
25.7532 | -100.3583
25.7569 | -100.3632
25.7606 | -100.3677
25.7643 | -100.3735
25.7679 | -100.3775
25.7716 | -100.383
25.7753 | -100.3904
25.779 | -100.3965
25.7826 | -100.4017
25.7863 | -100.4084
25.79 | -100.4151
25.7937 | -100.4211
25.7973 | -100.4271
25.801 | -100.4323
25.8047 | -100.4384
25.8083 | -100.4443
25.812 | -100.4525
25.8157 | -100.463
25.8194 | -100.4722
25.823 | -100.4794
25.8267 | -100.4855
25.8304 | -100.4934
25.8341 | -100.5007
25.8377 | -100.5086
25.8414 | -100.5165
25.8451 | -100.5253
25.8488 | -100.5367
25.8524 | -100.5481
25.8561 | -100.5583
25.8598 | -100.5681
25.8634 | -100.5806
25.8671 | -100.5901
25.8708 | -100.5982
25.8745 | -100.6077
25.8781 | -100.6146
25.8818 | -100.6215
25.8855 | -100.6302
25.8892 | -100.6422
25.8928 | -100.6537
25.8965 | -100.6626
25.9002 | -100.6684
25.9039 | -100.6723
25.9075 | -100.6772
25.9112 | -100.6833
25.9149 | -100.6903
25.9185 | -100.6988
25.9222 | -100.7053
25.9259 | -100.7111
25.9296 | -100.7168
25.9332 | -100.7218
25.9369 | -100.7252
25.9406 | -100.7271
25.9443 | -100.7311
25.9479 | -100.7333
25.9516 | -100.7393
25.9553 | -100.7454
25.959 | -100.7491
25.9626 | -100.752
25.9663 | -100.7525
25.97 | -100.7511
25.9736 | -100.7509
25.9773 | -100.7511
25.981 | -100.7488
25.9847 | -100.7472
25.9883 | -100.7458
25.992 | -100.7432
25.9957 | -100.7415
25.9994 | -100.7443
26.003 | -100.7452
26.0067 | -100.744
26.0104 | -100.7405
26.0141 | -100.734
26.0177 | -100.7282
26.0214 | -100.7249
26.0251 | -100.7218
26.0287 | -100.7173
26.0324 | -100.7144
26.0361 | -100.7127
26.0398 | -100.7085
26.0434 | -100.7045
26.0471 | -100.701
26.0508 | -100.6946
26.0545 | -100.6902
26.0581 | -100.688
26.0618 | -100.6835
26.0655 | -100.6774
26.0691 | -100.6712
26.0728 | -100.6624
26.0765 | -100.6563
26.0802 | -100.6489
26.0838 | -100.6403
26.0875 | -100.6342
26.0912 | -100.6286
26.0949 | -100.6229
26.0985 | -100.6161
26.1022 | -100.6099
26.1059 | -100.6035
26.1096 | -100.5987
26.1132 | -100.5952
26.1169 | -100.5923
26.1206 | -100.5883
26.1242 | -100.583
26.1279 | -100.5779
26.1316 | -100.5727
26.1353 | -100.5696
26.1389 | -100.5673
26.1426 | -100.5632
26.1463 | -100.5598
26.15 | -100.5568
26.1536 | -100.5528
26.1573 | -100.5494
26.161 | -100.5468
26.1647 | -100.5449
26.1683 | -100.5448
26.172 | -100.5423
26.1757 | -100.5394
26.1793 | -100.5378
26.183 | -100.5368
26.1867 | -100.5367
26.1904 | -100.5383
26.194 | -100.5426
26.1977 | -100.5453
26.2014 | -100.5499
26.2051 | -100.5542
26.2087 | -100.5569
26.2124 | -100.5602
26.2161 | -100.565
26.2198 | -100.5699
26.2234 | -100.5756
26.2271 | -100.5814
26.2308 | -100.5867
26.2344 | -100.5951
26.2381 | -100.6057
26.2418 | -100.6153
26.2455 | -100.6236
26.2491 | -100.634
26.2528 | -100.6437
26.2565 | -100.6519
26.2602 | -100.6614
26.2638 | -100.673
26.2675 | -100.6875
26.2712 | -100.7023
26.2749 | -100.7158
26.2785 | -100.7283
26.2822 | -100.7405
26.2859 | -100.7524
26.2895 | -100.7652
26.2932 | -100.7783
26.2969 | -100.7909
26.3006 | -100.8036
26.3042 | -100.8177
26.3079 | -100.8327
26.3116 | -100.8456
26.3153 | -100.8601
26.3189 | -100.8727
26.3226 | -100.8833
26.3263 | -100.8939
26.33 | -100.9041
26.3336 | -100.9151
26.3373 | -100.9272
26.341 | -100.941
26.3446 | -100.9508
26.3483 | -100.9605
26.352 | -100.9705
26.3557 | -100.9783
26.3593 | -100.9869
26.363 | -100.9973
26.3667 | -101.004
26.3704 | -101.008
26.374 | -101.0119
26.3777 | -101.0147
26.3814 | -101.0196
26.385 | -101.0245
26.3887 | -101.0273
26.3924 | -101.0283
26.3961 | -101.0256
26.3997 | -101.0235
26.4034 | -101.0219
26.4071 | -101.0223
26.4108 | -101.0216
26.4144 | -101.0205
26.4181 | -101.0177
26.4218 | -101.0124
26.4255 | -101.0053
26.4291 | -101.001
26.4328 | -100.9949
26.4365 | -100.9893
26.4401 | -100.9841
26.4438 | -100.9781
26.4475 | -100.9697
26.4512 | -100.9614
26.4548 | -100.9519
26.4585 | -100.9388
26.4622 | -100.9242
26.4659 | -100.9097
26.4695 | -100.8971
26.4732 | -100.8857
26.4769 | -100.8708
26.4806 | -100.8539
26.4842 | -100.839
26.4879 | -100.8233
26.4916 | -100.8056
26.4952 | -100.7881
26.4989 | -100.772
26.5026 | -100.7566
26.5063 | -100.7407
26.5099 | -100.7205
26.5136 | -100.702
26.5173 | -100.6857
26.521 | -100.6702
26.5246 | -100.6549
26.5283 | -100.638
26.532 | -100.6197
26.5357 | -100.6027
26.5393 | -100.5865
26.543 | -100.57
26.5467 | -100.5562
26.5503 | -100.5426
26.554 | -100.5268
26.5577 | -100.5097
26.5614 | -100.494
26.565 | -100.4771
26.5687 | -100.4603
26.5724 | -100.4458
26.5761 | -100.429
26.5797 | -100.413
26.5834 | -100.4005
26.5871 | -100.3888
26.5908 | -100.377
26.5944 | -100.3664
26.5981 | -100.3548
26.6018 | -100.3452
26.6054 | -100.3377
26.6091 | -100.3286
26.6128 | -100.3187
26.6165 | -100.3099
26.6201 | -100.3027
26.6238 | -100.2956
26.6275 | -100.2904
26.6312 | -100.2854
26.6348 | -100.2787
26.6385 | -100.273
26.6422 | -100.2694
26.6459 | -100.2671
26.6495 | -100.2659
26.6532 | -100.2643
26.6569 | -100.2627
26.6605 | -100.2618
26.6642 | -100.2615
26.6679 | -100.2617
26.6716 | -100.2629
26.6752 | -100.2645
26.6789 | -100.265
26.6826 | -100.2659
26.6863 | -100.2677
26.6899 | -100.2675
26.6936 | -100.2681
26.6973 | -100.2718
26.7009 | -100.2749
26.7046 | -100.2801
26.7083 | -100.2845
26.712 | -100.289
26.7156 | -100.2937
26.7193 | -100.2994
26.723 | -100.3056
26.7267 | -100.3127
26.7303 | -100.3193
26.734 | -100.3224
26.7377 | -100.3235
26.7414 | -100.3274
26.745 | -100.3332
26.7487 | -100.3391
26.7524 | -100.3455
26.756 | -100.3492
26.7597 | -100.354
26.7634 | -100.359
26.7671 | -100.364
26.7707 | -100.3701
26.7744 | -100.3746
26.7781 | -100.3759
26.7818 | -100.3751
26.7854 | -100.3775
26.7891 | -100.3785
26.7928 | -100.3763
26.7965 | -100.3772
26.8001 | -100.3779
26.8038 | -100.373
26.8075 | -100.369
26.8111 | -100.3656
26.8148 | -100.3624
26.8185 | -100.3602
26.8222 | -100.3575
26.8258 | -100.3537
26.8295 | -100.3498
26.8332 | -100.348
26.8369 | -100.3444
26.8405 | -100.3405
26.8442 | -100.3362
26.8479 | -100.331
26.8516 | -100.3272
26.8552 | -100.323
26.8589 | -100.3142
26.8626 | -100.3037
26.8662 | -100.2982
26.8699 | -100.2903
26.8736 | -100.2815
26.8773 | -100.2753
26.8809 | -100.2686
26.8846 | -100.261
26.8883 | -100.2572
26.892 | -100.2506
26.8956 | -100.2411
26.8993 | -100.2348
26.903 | -100.2279
26.9067 | -100.2211
26.9103 | -100.2173
26.914 | -100.2124
26.9177 | -100.2055
26.9213 | -100.201
26.925 | -100.1969
26.9287 | -100.1928
26.9324 | -100.1905
26.936 | -100.1883
26.9397 | -100.1827
26.9434 | -100.18
26.9471 | -100.1777
26.9507 | -100.1743
26.9544 | -100.1752
26.9581 | -100.1764
26.9618 | -100.1751
26.9654 | -100.1743
26.9691 | -100.1751
26.9728 | -100.1749
26.9764 | -100.1763
26.9801 | -100.1795
26.9838 | -100.1802
26.9875 | -100.183
26.9911 | -100.1896
26.9948 | -100.1954
26.9985 | -100.2004
27.0022 | -100.2057
27.0058 | -100.2077
27.0095 | -100.2108
27.0132 | -100.2175
27.0169 | -100.2216
27.0205 | -100.2243
27.0242 | -100.2324
27.0279 | -100.2378
27.0315 | -100.2432
27.0352 | -100.2485
27.0389 | -100.2555
27.0426 | -100.2623
27.0462 | -100.2702
27.0499 | -100.2776
27.0536 | -100.2849
27.0573 | -100.2908
27.0609 | -100.2975
27.0646 | -100.3053
27.0683 | -100.3111
27.072 | -100.3166
27.0756 | -100.3239
27.0793 | -100.3282
27.083 | -100.3342
27.0866 | -100.341
27.0903 | -100.3467
27.094 | -100.3522
27.0977 | -100.3587
27.1013 | -100.3653
27.105 | -100.3704
27.1087 | -100.3756
27.1124 | -100.3817
27.116 | -100.3881
27.1197 | -100.3958
27.1234 | -100.4022
27.127 | -100.4056
27.1307 | -100.4103
27.1344 | -100.4154
27.1381 | -100.4206
27.1417 | -100.4268
27.1454 | -100.4317
27.1491 | -100.435
27.1528 | -100.4394
27.1564 | -100.4454
27.1601 | -100.4536
27.1638 | -100.4625
27.1675 | -100.4692
27.1711 | -100.4739
27.1748 | -100.4816
27.1785 | -100.4891
27.1821 | -100.4967
27.1858 | -100.505
27.1895 | -100.5104
27.1932 | -100.515
27.1968 | -100.5196
27.2005 | -100.5235
27.2042 | -100.5267
27.2079 | -100.5316
27.2115 | -100.5368
27.2152 | -100.541
27.2189 | -100.5448
27.2226 | -100.5475
27.2262 | -100.5497
27.2299 | -100.5507
27.2336 | -100.5518
27.2372 | -100.5531
27.2409 | -100.557
27.2446 | -100.5626
27.2483 | -100.5647
27.2519 | -100.5669
27.2556 | -100.5685
27.2593 | -100.5698
27.263 | -100.5712
27.2666 | -100.5728
27.2703 | -100.5733
27.274 | -100.5741
27.2777 | -100.5755
27.2813 | -100.5744
27.285 | -100.5745
27.2887 | -100.5754
27.2923 | -100.5739
27.296 | -100.5728
27.2997 | -100.5729
27.3034 | -100.5702
27.307 | -100.567
27.3107 | -100.5637
27.3144 | -100.5606
27.3181 | -100.5594
27.3217 | -100.5607
27.3254 | -100.5614
27.3291 | -100.5608
27.3328 | -100.56
27.3364 | -100.5567
27.3401 | -100.5537
27.3438 | -100.5519
27.3474 | -100.5493
27.3511 | -100.5463
27.3548 | -100.5441
27.3585 | -100.542
27.3621 | -100.5372
27.3658 | -100.5349
27.3695 | -100.534
27.3732 | -100.531
27.3768 | -100.5275
27.3805 | -100.5242
27.3842 | -100.5211
27.3879 | -100.5207
27.3915 | -100.5209
27.3952 | -100.5204
27.3989 | -100.5214
27.4025 | -100.5219
27.4062 | -100.5201
27.4099 | -100.5173
27.4136 | -100.5166
27.4172 | -100.5155
27.4209 | -100.5157
27.4246 | -100.5135
27.4283 | -100.5128
27.4319 | -100.5117
27.4356 | -100.5104
27.4393 | -100.5093
27.4429 | -100.5094
27.4466 | -100.51
27.4503 | -100.5095
27.454 | -100.5108
27.4576 | -100.5133
27.4613 | -100.5125
27.465 | -100.5134
27.4687 | -100.5146
27.4723 | -100.5139
27.476 | -100.5134
27.4797 | -100.5132
27.4834 | -100.5116
27.487 | -100.5115
27.4907 | -100.5105
27.4944 | -100.5113
27.498 | -100.5119
27.5017 | -100.513
27.5054 | -100.5145
27.5091 | -100.5157
27.5127 | -100.5179
27.5164 | -100.518
27.5201 | -100.5157
27.5238 | -100.5163
27.5274 | -100.5179
27.5311 | -100.5197
27.5348 | -100.5236
27.5385 | -100.5271
27.5421 | -100.5285
27.5458 | -100.5283
27.5495 | -100.5291
27.5531 | -100.5349
27.5568 | -100.5372
27.5605 | -100.5428
27.5642 | -100.5475
27.5678 | -100.5489
27.5715 | -100.5501
27.5752 | -100.5528
27.5789 | -100.5569
27.5825 | -100.5613
27.5862 | -100.5663
27.5899 | -100.572
27.5936 | -100.5774
27.5972 | -100.5813
27.6009 | -100.5838
27.6046 | -100.5855
27.6082 | -100.5874
27.6119 | -100.5895
27.6156 | -100.5913
27.6193 | -100.5962
27.6229 | -100.5999
27.6266 | -100.6037
27.6303 | -100.6044
27.634 | -100.6023
27.6376 | -100.5991
27.6413 | -100.5955
27.645 | -100.5935
27.6487 | -100.5905
27.6523 | -100.5869
27.656 | -100.5822
27.6597 | -100.5776
27.6633 | -100.5718
27.667 | -100.5649
27.6707 | -100.5584
27.6744 | -100.5522
27.678 | -100.5457
27.6817 | -100.5385
27.6854 | -100.5305
27.6891 | -100.5218
27.6927 | -100.5123
27.6964 | -100.5019
27.7001 | -100.4912
27.7038 | -100.4786
27.7074 | -100.4646
27.7111 | -100.4493
27.7148 | -100.4329
27.7184 | -100.4181
27.7221 | -100.4029
27.7258 | -100.3871
27.7295 | -100.372
27.7331 | -100.3554
27.7368 | -100.3373
27.7405 | -100.3183
27.7442 | -100.3001
27.7478 | -100.2826
27.7515 | -100.2664
27.7552 | -100.2493
27.7588 | -100.2306
27.7625 | -100.2118
27.7662 | -100.1934
27.7699 | -100.1751
27.7735 | -100.1573
27.7772 | -100.1411
27.7809 | -100.1244
27.7846 | -100.1075
27.7882 | -100.0925
27.7919 | -100.0771
27.7956 | -100.0617
27.7993 | -100.0489
27.8029 | -100.0343
27.8066 | -100.021
27.8103 | -100.009
27.8139 | -99.9995
27.8176 | -99.9913
27.8213 | -99.984
27.825 | -99.9783
27.8286 | -99.9728
27.8323 | -99.9681
27.836 | -99.9648
27.8397 | -99.9617
27.8433 | -99.9601
27.847 | -99.9613
27.8507 | -99.964
27.8544 | -99.9661
27.858 | -99.9692
27.8617 | -99.9738
27.8654 | -99.9798
27.869 | -99.9877
27.8727 | -99.9975
27.8764 | -100.0069
27.8801 | -100.0191
27.8837 | -100.0313
27.8874 | -100.0443
27.8911 | -100.0584
27.8948 | -100.0727
27.8984 | -100.0879
27.9021 | -100.1039
27.9058 | -100.1206
27.9095 | -100.1378
27.9131 | -100.1553
27.9168 | -100.1733
27.9205 | -100.1934
27.9241 | -100.2125
27.9278 | -100.2322
27.9315 | -100.2534
27.9352 | -100.2757
27.9388 | -100.2968
27.9425 | -100.3187
27.9462 | -100.3412
27.9499 | -100.3634
27.9535 | -100.3855
27.9572 | -100.4093
27.9609 | -100.4333
27.9646 | -100.4561
27.9682 | -100.48
27.9719 | -100.5044
27.9756 | -100.5262
27.9792 | -100.5461
27.9829 | -100.5627
27.9866 | -100.5799
27.9903 | -100.6011
27.9939 | -100.6226
27.9976 | -100.6438
28.0013 | -100.6648