## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 229 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 224
Hood | 02-92B/P | 163
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1014.2 | hPa |  | PRC PRO 01
Humidity | 45.2 | %/Hr
Internal temperature of calibration bench | 23.3 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1011 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.1 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.06 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.02 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3117 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 16/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 24/05/2018
Engineer | Anthony HUE | Approval |  | 29/05/2018
CEO | Robert GLEMOT | Approval |  | 29/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3387 | 57.9113
0.012209 | -35.986 | 38.9503
0.018313 | -34.9763 | 28.5144
0.024417 | -34.5105 | 22.0757
0.030521 | -34.3335 | 18.1242
0.036626 | -34.1837 | 15.0863
0.04273 | -34.0898 | 12.8808
0.048834 | -34.0381 | 11.2493
0.054938 | -33.9995 | 10.0111
0.061043 | -33.9663 | 9.0014
0.067147 | -33.952 | 8.1925
0.073251 | -33.9412 | 7.4443
0.079355 | -33.9361 | 6.7996
0.08546 | -33.931 | 6.249
0.091564 | -33.9242 | 5.7936
0.097668 | -33.9175 | 5.4221
0.10377 | -33.9115 | 5.1072
0.10988 | -33.9059 | 4.7962
0.11598 | -33.9017 | 4.5057
0.12209 | -33.8999 | 4.2338
0.12819 | -33.9002 | 3.9769
0.13429 | -33.901 | 3.7353
0.1404 | -33.9001 | 3.5285
0.1465 | -33.8993 | 3.3309
0.15261 | -33.8981 | 3.1569
0.15871 | -33.8971 | 3.0089
0.16482 | -33.896 | 2.8814
0.17092 | -33.8956 | 2.772
0.17702 | -33.8951 | 2.6563
0.18313 | -33.8948 | 2.5436
0.18923 | -33.8946 | 2.4305
0.19534 | -33.8943 | 2.3196
0.20144 | -33.8942 | 2.2112
0.20754 | -33.8941 | 2.1163
0.21365 | -33.894 | 2.0216
0.21975 | -33.894 | 1.9335
0.22586 | -33.8943 | 1.8509
0.23196 | -33.8948 | 1.7735
0.23807 | -33.8956 | 1.6995
0.24417 | -33.8967 | 1.6301
0.25027 | -33.8979 | 1.5599
0.25638 | -33.8988 | 1.4922
0.26248 | -33.8995 | 1.4263
0.26859 | -33.9001 | 1.3638
0.27469 | -33.9006 | 1.3026
0.2808 | -33.9011 | 1.2462
0.2869 | -33.9017 | 1.1884
0.293 | -33.9023 | 1.1329
0.29911 | -33.903 | 1.0801
0.30521 | -33.9037 | 1.0277
0.31132 | -33.9045 | 0.98055
0.31742 | -33.9053 | 0.93295
0.32353 | -33.906 | 0.88707
0.32963 | -33.9067 | 0.84391
0.33573 | -33.9073 | 0.8036
0.34184 | -33.9078 | 0.76497
0.34794 | -33.9081 | 0.73085
0.35405 | -33.9086 | 0.69438
0.36015 | -33.9091 | 0.66016
0.36626 | -33.9094 | 0.62613
0.37236 | -33.9097 | 0.59268
0.37846 | -33.91 | 0.55795
0.38457 | -33.9101 | 0.52127
0.39067 | -33.9103 | 0.4865
0.39678 | -33.9104 | 0.45131
0.40288 | -33.9104 | 0.41592
0.40899 | -33.9104 | 0.38034
0.41509 | -33.9103 | 0.34581
0.42119 | -33.9102 | 0.3137
0.4273 | -33.9102 | 0.28056
0.4334 | -33.9102 | 0.24827
0.43951 | -33.9104 | 0.21615
0.44561 | -33.9107 | 0.18573
0.45172 | -33.9114 | 0.15764
0.45782 | -33.912 | 0.12808
0.46392 | -33.9126 | 0.098637
0.47003 | -33.9134 | 0.070081
0.47613 | -33.9142 | 0.041548
0.48224 | -33.915 | 0.012976
0.48834 | -33.9157 | -0.014805
0.49445 | -33.9165 | -0.042975
0.50055 | -33.9172 | -0.070404
0.50665 | -33.9179 | -0.097422
0.51276 | -33.9184 | -0.12498
0.51886 | -33.9189 | -0.15118
0.52497 | -33.9192 | -0.17549
0.53107 | -33.9196 | -0.20106
0.53717 | -33.92 | -0.22562
0.54328 | -33.9202 | -0.24963
0.54938 | -33.9206 | -0.27273
0.55549 | -33.9209 | -0.29503
0.56159 | -33.9213 | -0.3152
0.5677 | -33.9216 | -0.3366
0.5738 | -33.9221 | -0.35824
0.5799 | -33.9225 | -0.37941
0.58601 | -33.923 | -0.39953
0.59211 | -33.9236 | -0.41762
0.59822 | -33.924 | -0.437
0.60432 | -33.9245 | -0.45723
0.61043 | -33.925 | -0.47623
0.61653 | -33.9255 | -0.49588
0.62263 | -33.9259 | -0.51544
0.62874 | -33.9262 | -0.53591
0.63484 | -33.9265 | -0.55545
0.64095 | -33.9268 | -0.57534
0.64705 | -33.927 | -0.59548
0.65316 | -33.9272 | -0.61586
0.65926 | -33.9273 | -0.63709
0.66536 | -33.9272 | -0.65766
0.67147 | -33.9271 | -0.67756
0.67757 | -33.9271 | -0.69768
0.68368 | -33.9269 | -0.71815
0.68978 | -33.9267 | -0.73811
0.69589 | -33.9265 | -0.75802
0.70199 | -33.9264 | -0.77736
0.70809 | -33.9262 | -0.79687
0.7142 | -33.9261 | -0.81609
0.7203 | -33.9259 | -0.83465
0.72641 | -33.9258 | -0.85323
0.73251 | -33.9257 | -0.87139
0.73862 | -33.926 | -0.88822
0.74472 | -33.926 | -0.90558
0.75082 | -33.9261 | -0.92325
0.75693 | -33.9263 | -0.94034
0.76303 | -33.9266 | -0.95732
0.76914 | -33.9272 | -0.97301
0.77524 | -33.9276 | -0.98936
0.78135 | -33.9281 | -1.0055
0.78745 | -33.9286 | -1.0215
0.79355 | -33.9291 | -1.0374
0.79966 | -33.9297 | -1.053
0.80576 | -33.9306 | -1.0685
0.81187 | -33.9312 | -1.0843
0.81797 | -33.9319 | -1.0998
0.82408 | -33.9325 | -1.1153
0.83018 | -33.9332 | -1.1306
0.83628 | -33.9339 | -1.1458
0.84239 | -33.9345 | -1.1614
0.84849 | -33.9351 | -1.177
0.8546 | -33.9357 | -1.1925
0.8607 | -33.9363 | -1.208
0.86681 | -33.9369 | -1.2232
0.87291 | -33.9374 | -1.2383
0.87901 | -33.9378 | -1.2533
0.88512 | -33.9383 | -1.2683
0.89122 | -33.9387 | -1.2833
0.89733 | -33.9391 | -1.2978
0.90343 | -33.9395 | -1.3127
0.90953 | -33.9398 | -1.327
0.91564 | -33.9401 | -1.3416
0.92174 | -33.9404 | -1.3565
0.92785 | -33.9407 | -1.3706
0.93395 | -33.941 | -1.3846
0.94006 | -33.9413 | -1.3985
0.94616 | -33.9415 | -1.4122
0.95226 | -33.9417 | -1.4258
0.95837 | -33.942 | -1.4397
0.96447 | -33.9423 | -1.4539
0.97058 | -33.9425 | -1.4681
0.97668 | -33.9428 | -1.4819
0.98279 | -33.9429 | -1.4957
0.98889 | -33.9432 | -1.5093
0.99499 | -33.9434 | -1.5232
1.0011 | -33.9436 | -1.5366
1.0072 | -33.9437 | -1.5502
1.0133 | -33.9439 | -1.5635
1.0194 | -33.944 | -1.5772
1.0255 | -33.9442 | -1.591
1.0316 | -33.9443 | -1.6045
1.0377 | -33.9445 | -1.6186
1.0438 | -33.9446 | -1.6324
1.0499 | -33.9446 | -1.6465
1.056 | -33.9447 | -1.6604
1.0621 | -33.9448 | -1.6744
1.0682 | -33.9449 | -1.688
1.0743 | -33.945 | -1.702
1.0805 | -33.9451 | -1.7159
1.0866 | -33.9451 | -1.73
1.0927 | -33.9452 | -1.744
1.0988 | -33.9452 | -1.7581
1.1049 | -33.9453 | -1.772
1.111 | -33.9453 | -1.7861
1.1171 | -33.9454 | -1.8
1.1232 | -33.9455 | -1.814
1.1293 | -33.9456 | -1.8279
1.1354 | -33.9458 | -1.8417
1.1415 | -33.9459 | -1.8557
1.1476 | -33.946 | -1.8697
1.1537 | -33.9461 | -1.8835
1.1598 | -33.9463 | -1.8971
1.1659 | -33.9464 | -1.9104
1.172 | -33.9466 | -1.9242
1.1781 | -33.9468 | -1.9377
1.1842 | -33.947 | -1.9512
1.1903 | -33.9473 | -1.9644
1.1964 | -33.9475 | -1.9782
1.2025 | -33.9477 | -1.9915
1.2086 | -33.9479 | -2.0047
1.2147 | -33.9481 | -2.0182
1.2209 | -33.9484 | -2.0313
1.227 | -33.9486 | -2.0443
1.2331 | -33.9489 | -2.0574
1.2392 | -33.9491 | -2.0705
1.2453 | -33.9494 | -2.0831
1.2514 | -33.9496 | -2.0959
1.2575 | -33.9498 | -2.1089
1.2636 | -33.95 | -2.1217
1.2697 | -33.9503 | -2.1344
1.2758 | -33.9505 | -2.1475
1.2819 | -33.9507 | -2.1603
1.288 | -33.9509 | -2.173
1.2941 | -33.9512 | -2.1852
1.3002 | -33.9514 | -2.1977
1.3063 | -33.9515 | -2.2102
1.3124 | -33.9517 | -2.2226
1.3185 | -33.9519 | -2.2352
1.3246 | -33.9521 | -2.2475
1.3307 | -33.9522 | -2.2599
1.3368 | -33.9524 | -2.2721
1.3429 | -33.9525 | -2.2847
1.349 | -33.9526 | -2.2972
1.3551 | -33.9528 | -2.3098
1.3613 | -33.9529 | -2.3221
1.3674 | -33.9531 | -2.3343
1.3735 | -33.9532 | -2.3468
1.3796 | -33.9533 | -2.3591
1.3857 | -33.9534 | -2.3713
1.3918 | -33.9536 | -2.3833
1.3979 | -33.9537 | -2.3956
1.404 | -33.9539 | -2.4076
1.4101 | -33.9541 | -2.4198
1.4162 | -33.9543 | -2.4321
1.4223 | -33.9544 | -2.4443
1.4284 | -33.9547 | -2.4564
1.4345 | -33.9548 | -2.4689
1.4406 | -33.9551 | -2.4812
1.4467 | -33.9553 | -2.4931
1.4528 | -33.9555 | -2.5054
1.4589 | -33.9558 | -2.5178
1.465 | -33.9559 | -2.5298
1.4711 | -33.9561 | -2.542
1.4772 | -33.9565 | -2.5542
1.4833 | -33.9568 | -2.5663
1.4894 | -33.9571 | -2.5786
1.4955 | -33.9574 | -2.5905
1.5016 | -33.9577 | -2.6025
1.5078 | -33.958 | -2.6147
1.5139 | -33.9583 | -2.6267
1.52 | -33.9586 | -2.6386
1.5261 | -33.9588 | -2.6506
1.5322 | -33.9591 | -2.663
1.5383 | -33.9594 | -2.6749
1.5444 | -33.9598 | -2.6866
1.5505 | -33.9601 | -2.6988
1.5566 | -33.9604 | -2.7107
1.5627 | -33.9607 | -2.7227
1.5688 | -33.9611 | -2.7349
1.5749 | -33.9613 | -2.7468
1.581 | -33.9616 | -2.7586
1.5871 | -33.9619 | -2.7707
1.5932 | -33.9622 | -2.7827
1.5993 | -33.9625 | -2.7949
1.6054 | -33.9628 | -2.8069
1.6115 | -33.9631 | -2.8188
1.6176 | -33.9633 | -2.8307
1.6237 | -33.9636 | -2.8426
1.6298 | -33.9639 | -2.8543
1.6359 | -33.9641 | -2.866
1.642 | -33.9644 | -2.8777
1.6482 | -33.9647 | -2.8895
1.6543 | -33.965 | -2.9014
1.6604 | -33.9653 | -2.9131
1.6665 | -33.9655 | -2.9249
1.6726 | -33.9658 | -2.9366
1.6787 | -33.9661 | -2.9481
1.6848 | -33.9663 | -2.9599
1.6909 | -33.9666 | -2.9718
1.697 | -33.9668 | -2.9835
1.7031 | -33.9671 | -2.9949
1.7092 | -33.9672 | -3.0064
1.7153 | -33.9675 | -3.0177
1.7214 | -33.9676 | -3.0293
1.7275 | -33.9679 | -3.0411
1.7336 | -33.9681 | -3.0527
1.7397 | -33.9684 | -3.0644
1.7458 | -33.9686 | -3.0762
1.7519 | -33.9688 | -3.0879
1.758 | -33.969 | -3.0991
1.7641 | -33.9693 | -3.1108
1.7702 | -33.9695 | -3.1229
1.7763 | -33.9697 | -3.1347
1.7824 | -33.9699 | -3.1465
1.7885 | -33.9701 | -3.1588
1.7947 | -33.9704 | -3.1705
1.8008 | -33.9706 | -3.1819
1.8069 | -33.9709 | -3.1936
1.813 | -33.9711 | -3.2056
1.8191 | -33.9713 | -3.2175
1.8252 | -33.9715 | -3.2293
1.8313 | -33.9718 | -3.2412
1.8374 | -33.972 | -3.2529
1.8435 | -33.9722 | -3.2647
1.8496 | -33.9724 | -3.2769
1.8557 | -33.9727 | -3.2887
1.8618 | -33.9729 | -3.3005
1.8679 | -33.9731 | -3.3125
1.874 | -33.9734 | -3.3243
1.8801 | -33.9736 | -3.3362
1.8862 | -33.9738 | -3.3479
1.8923 | -33.974 | -3.3598
1.8984 | -33.9743 | -3.3718
1.9045 | -33.9745 | -3.3836
1.9106 | -33.9747 | -3.3957
1.9167 | -33.9749 | -3.4074
1.9228 | -33.9752 | -3.4191
1.9289 | -33.9754 | -3.431
1.9351 | -33.9757 | -3.4426
1.9412 | -33.9759 | -3.4541
1.9473 | -33.9761 | -3.4658
1.9534 | -33.9764 | -3.4776
1.9595 | -33.9766 | -3.4892
1.9656 | -33.9768 | -3.501
1.9717 | -33.9771 | -3.5126
1.9778 | -33.9773 | -3.5244
1.9839 | -33.9776 | -3.5359
1.99 | -33.9778 | -3.5477
1.9961 | -33.978 | -3.5592
2.0022 | -33.9782 | -3.5704
2.0083 | -33.9784 | -3.5819
2.0144 | -33.9787 | -3.5935
2.0205 | -33.979 | -3.6049
2.0266 | -33.9792 | -3.6163
2.0327 | -33.9795 | -3.6279
2.0388 | -33.9798 | -3.639
2.0449 | -33.98 | -3.6502
2.051 | -33.9802 | -3.6616
2.0571 | -33.9804 | -3.673
2.0632 | -33.9806 | -3.6843
2.0693 | -33.9809 | -3.696
2.0754 | -33.9811 | -3.7073
2.0816 | -33.9814 | -3.7185
2.0877 | -33.9816 | -3.7295
2.0938 | -33.9819 | -3.741
2.0999 | -33.9822 | -3.7524
2.106 | -33.9824 | -3.7634
2.1121 | -33.9826 | -3.7747
2.1182 | -33.9829 | -3.7859
2.1243 | -33.9831 | -3.797
2.1304 | -33.9833 | -3.8082
2.1365 | -33.9836 | -3.8196
2.1426 | -33.9838 | -3.8309
2.1487 | -33.984 | -3.8422
2.1548 | -33.9843 | -3.8537
2.1609 | -33.9846 | -3.8646
2.167 | -33.9848 | -3.8759
2.1731 | -33.985 | -3.8874
2.1792 | -33.9852 | -3.8986
2.1853 | -33.9855 | -3.9099
2.1914 | -33.9857 | -3.9212
2.1975 | -33.986 | -3.9327
2.2036 | -33.9862 | -3.9444
2.2097 | -33.9864 | -3.9557
2.2158 | -33.9867 | -3.9671
2.222 | -33.9869 | -3.9783
2.2281 | -33.9871 | -3.9896
2.2342 | -33.9874 | -4.0011
2.2403 | -33.9876 | -4.0125
2.2464 | -33.9878 | -4.0241
2.2525 | -33.9881 | -4.0356
2.2586 | -33.9883 | -4.047
2.2647 | -33.9885 | -4.0583
2.2708 | -33.9887 | -4.0696
2.2769 | -33.989 | -4.0811
2.283 | -33.9892 | -4.093
2.2891 | -33.9894 | -4.1046
2.2952 | -33.9896 | -4.1161
2.3013 | -33.9898 | -4.1277
2.3074 | -33.99 | -4.1392
2.3135 | -33.9902 | -4.1511
2.3196 | -33.9904 | -4.1632
2.3257 | -33.9906 | -4.1751
2.3318 | -33.9908 | -4.1872
2.3379 | -33.991 | -4.199
2.344 | -33.9912 | -4.211
2.3501 | -33.9913 | -4.2226
2.3562 | -33.9916 | -4.2346
2.3623 | -33.9918 | -4.2465
2.3685 | -33.992 | -4.2582
2.3746 | -33.9922 | -4.2702
2.3807 | -33.9924 | -4.2821
2.3868 | -33.9926 | -4.2941
2.3929 | -33.9927 | -4.3061
2.399 | -33.9929 | -4.318
2.4051 | -33.9931 | -4.3297
2.4112 | -33.9934 | -4.3415
2.4173 | -33.9936 | -4.3535
2.4234 | -33.9938 | -4.3656
2.4295 | -33.994 | -4.3775
2.4356 | -33.9942 | -4.3897
2.4417 | -33.9944 | -4.4018
2.4478 | -33.9946 | -4.4139
2.4539 | -33.9948 | -4.4258
2.46 | -33.995 | -4.4378
2.4661 | -33.9952 | -4.4497
2.4722 | -33.9954 | -4.4616
2.4783 | -33.9956 | -4.4735
2.4844 | -33.9958 | -4.4853
2.4905 | -33.996 | -4.4972
2.4966 | -33.9962 | -4.5088
2.5027 | -33.9964 | -4.5205
2.5089 | -33.9966 | -4.5324
2.515 | -33.9968 | -4.5445
2.5211 | -33.997 | -4.5561
2.5272 | -33.9972 | -4.5676
2.5333 | -33.9974 | -4.5791
2.5394 | -33.9976 | -4.5905
2.5455 | -33.9978 | -4.6024
2.5516 | -33.998 | -4.6139
2.5577 | -33.9983 | -4.6255
2.5638 | -33.9985 | -4.6369
2.5699 | -33.9986 | -4.6485
2.576 | -33.9989 | -4.6604
2.5821 | -33.999 | -4.6718
2.5882 | -33.9992 | -4.6832
2.5943 | -33.9995 | -4.6948
2.6004 | -33.9997 | -4.7059
2.6065 | -33.9999 | -4.7174
2.6126 | -34.0001 | -4.7288
2.6187 | -34.0003 | -4.7399
2.6248 | -34.0005 | -4.7512
2.6309 | -34.0007 | -4.7626
2.637 | -34.0009 | -4.7734
2.6431 | -34.0011 | -4.7846
2.6492 | -34.0014 | -4.796
2.6554 | -34.0016 | -4.8071
2.6615 | -34.0018 | -4.8183
2.6676 | -34.002 | -4.8294
2.6737 | -34.0022 | -4.8401
2.6798 | -34.0024 | -4.8511
2.6859 | -34.0026 | -4.8621
2.692 | -34.0029 | -4.8731
2.6981 | -34.0031 | -4.8842
2.7042 | -34.0033 | -4.8948
2.7103 | -34.0035 | -4.9056
2.7164 | -34.0037 | -4.9164
2.7225 | -34.0039 | -4.9274
2.7286 | -34.0041 | -4.9382
2.7347 | -34.0043 | -4.9491
2.7408 | -34.0045 | -4.9595
2.7469 | -34.0047 | -4.9705
2.753 | -34.0049 | -4.9811
2.7591 | -34.0051 | -4.9919
2.7652 | -34.0053 | -5.0027
2.7713 | -34.0055 | -5.0134
2.7774 | -34.0057 | -5.0239
2.7835 | -34.006 | -5.0345
2.7896 | -34.0061 | -5.045
2.7958 | -34.0063 | -5.0557
2.8019 | -34.0065 | -5.0663
2.808 | -34.0068 | -5.0768
2.8141 | -34.007 | -5.0871
2.8202 | -34.0072 | -5.0977
2.8263 | -34.0074 | -5.1082
2.8324 | -34.0076 | -5.119
2.8385 | -34.0079 | -5.1293
2.8446 | -34.0081 | -5.1395
2.8507 | -34.0083 | -5.1501
2.8568 | -34.0085 | -5.1607
2.8629 | -34.0087 | -5.1712
2.869 | -34.0089 | -5.1818
2.8751 | -34.0091 | -5.1923
2.8812 | -34.0093 | -5.2027
2.8873 | -34.0095 | -5.2132
2.8934 | -34.0098 | -5.2238
2.8995 | -34.01 | -5.2342
2.9056 | -34.0102 | -5.2448
2.9117 | -34.0104 | -5.2553
2.9178 | -34.0106 | -5.2656
2.9239 | -34.0108 | -5.2764
2.93 | -34.0111 | -5.2868
2.9361 | -34.0113 | -5.2973
2.9423 | -34.0115 | -5.3077
2.9484 | -34.0117 | -5.318
2.9545 | -34.012 | -5.3285
2.9606 | -34.0122 | -5.3389
2.9667 | -34.0124 | -5.3494
2.9728 | -34.0126 | -5.3601
2.9789 | -34.0128 | -5.3706
2.985 | -34.0131 | -5.3811
2.9911 | -34.0133 | -5.3916
2.9972 | -34.0135 | -5.402
3.0033 | -34.0137 | -5.4127
3.0094 | -34.0139 | -5.4233
3.0155 | -34.0141 | -5.4338
3.0216 | -34.0143 | -5.4443
3.0277 | -34.0145 | -5.4549
3.0338 | -34.0148 | -5.4655
3.0399 | -34.015 | -5.4759
3.046 | -34.0152 | -5.4865
3.0521 | -34.0154 | -5.4971
3.0582 | -34.0156 | -5.5079
3.0643 | -34.0158 | -5.5186
3.0704 | -34.016 | -5.5292
3.0765 | -34.0162 | -5.5397
3.0827 | -34.0164 | -5.5503
3.0888 | -34.0166 | -5.5609
3.0949 | -34.0169 | -5.5716
3.101 | -34.0171 | -5.5822
3.1071 | -34.0173 | -5.5928
3.1132 | -34.0175 | -5.6036
3.1193 | -34.0177 | -5.6142
3.1254 | -34.0179 | -5.6249
3.1315 | -34.0181 | -5.6356
3.1376 | -34.0183 | -5.6462
3.1437 | -34.0186 | -5.6569
3.1498 | -34.0188 | -5.6676
3.1559 | -34.019 | -5.6782
3.162 | -34.0192 | -5.6892
3.1681 | -34.0194 | -5.7001
3.1742 | -34.0197 | -5.7107
3.1803 | -34.0199 | -5.7216
3.1864 | -34.0201 | -5.7323
3.1925 | -34.0203 | -5.7429
3.1986 | -34.0205 | -5.7538
3.2047 | -34.0207 | -5.7647
3.2108 | -34.021 | -5.7754
3.2169 | -34.0212 | -5.7862
3.223 | -34.0214 | -5.7971
3.2292 | -34.0216 | -5.8081
3.2353 | -34.0218 | -5.8188
3.2414 | -34.022 | -5.8297
3.2475 | -34.0222 | -5.8404
3.2536 | -34.0224 | -5.8511
3.2597 | -34.0226 | -5.862
3.2658 | -34.0228 | -5.8726
3.2719 | -34.023 | -5.8834
3.278 | -34.0233 | -5.8941
3.2841 | -34.0235 | -5.905
3.2902 | -34.0237 | -5.9157
3.2963 | -34.0239 | -5.9265
3.3024 | -34.0241 | -5.9374
3.3085 | -34.0243 | -5.9482
3.3146 | -34.0245 | -5.9589
3.3207 | -34.0247 | -5.9697
3.3268 | -34.0249 | -5.9805
3.3329 | -34.0252 | -5.9913
3.339 | -34.0254 | -6.0021
3.3451 | -34.0256 | -6.0129
3.3512 | -34.0258 | -6.0238
3.3573 | -34.026 | -6.0347
3.3634 | -34.0262 | -6.0455
3.3696 | -34.0264 | -6.0564
3.3757 | -34.0266 | -6.0673
3.3818 | -34.0268 | -6.0781
3.3879 | -34.027 | -6.089
3.394 | -34.0272 | -6.0998
3.4001 | -34.0274 | -6.1106
3.4062 | -34.0276 | -6.1215
3.4123 | -34.0278 | -6.1323
3.4184 | -34.028 | -6.1432
3.4245 | -34.0282 | -6.154
3.4306 | -34.0284 | -6.1649
3.4367 | -34.0286 | -6.1758
3.4428 | -34.0288 | -6.1866
3.4489 | -34.029 | -6.1974
3.455 | -34.0292 | -6.2083
3.4611 | -34.0294 | -6.2191
3.4672 | -34.0296 | -6.2299
3.4733 | -34.0298 | -6.2408
3.4794 | -34.03 | -6.2516
3.4855 | -34.0301 | -6.2625
3.4916 | -34.0303 | -6.2732
3.4977 | -34.0305 | -6.2842
3.5038 | -34.0307 | -6.2952
3.5099 | -34.0309 | -6.3059
3.5161 | -34.0311 | -6.3167
3.5222 | -34.0313 | -6.3275
3.5283 | -34.0315 | -6.3383
3.5344 | -34.0317 | -6.3492
3.5405 | -34.0319 | -6.36
3.5466 | -34.0321 | -6.3709
3.5527 | -34.0323 | -6.3817
3.5588 | -34.0324 | -6.3926
3.5649 | -34.0326 | -6.4036
3.571 | -34.0328 | -6.4144
3.5771 | -34.033 | -6.4255
3.5832 | -34.0332 | -6.4363
3.5893 | -34.0334 | -6.4471
3.5954 | -34.0336 | -6.4579
3.6015 | -34.0338 | -6.4688
3.6076 | -34.034 | -6.4797
3.6137 | -34.0342 | -6.4906
3.6198 | -34.0344 | -6.5013
3.6259 | -34.0346 | -6.5123
3.632 | -34.0348 | -6.5231
3.6381 | -34.035 | -6.5339
3.6442 | -34.0352 | -6.5447
3.6503 | -34.0354 | -6.5554
3.6565 | -34.0356 | -6.5663
3.6626 | -34.0358 | -6.5772
3.6687 | -34.036 | -6.5881
3.6748 | -34.0362 | -6.5989
3.6809 | -34.0364 | -6.6098
3.687 | -34.0366 | -6.6208
3.6931 | -34.0368 | -6.6316
3.6992 | -34.037 | -6.6423
3.7053 | -34.0372 | -6.6531
3.7114 | -34.0374 | -6.664
3.7175 | -34.0376 | -6.6748
3.7236 | -34.0378 | -6.6855
3.7297 | -34.038 | -6.6965
3.7358 | -34.0382 | -6.7074
3.7419 | -34.0384 | -6.7183
3.748 | -34.0386 | -6.7293
3.7541 | -34.0389 | -6.7401
3.7602 | -34.0391 | -6.7508
3.7663 | -34.0393 | -6.7617
3.7724 | -34.0395 | -6.7725
3.7785 | -34.0397 | -6.7831
3.7846 | -34.0399 | -6.7941
3.7907 | -34.0402 | -6.8049
3.7969 | -34.0404 | -6.8158
3.803 | -34.0406 | -6.8266
3.8091 | -34.0408 | -6.8374
3.8152 | -34.0411 | -6.8481
3.8213 | -34.0413 | -6.859
3.8274 | -34.0415 | -6.8698
3.8335 | -34.0417 | -6.8806
3.8396 | -34.042 | -6.8914
3.8457 | -34.0422 | -6.9021
3.8518 | -34.0424 | -6.9129
3.8579 | -34.0427 | -6.9237
3.864 | -34.0429 | -6.9345
3.8701 | -34.0432 | -6.9454
3.8762 | -34.0434 | -6.9562
3.8823 | -34.0436 | -6.967
3.8884 | -34.0439 | -6.9778
3.8945 | -34.0441 | -6.9886
3.9006 | -34.0444 | -6.9994
3.9067 | -34.0446 | -7.0103
3.9128 | -34.0449 | -7.0211
3.9189 | -34.0451 | -7.0318
3.925 | -34.0453 | -7.0426
3.9311 | -34.0456 | -7.0533
3.9372 | -34.0458 | -7.0641
3.9434 | -34.0461 | -7.0749
3.9495 | -34.0463 | -7.0858
3.9556 | -34.0466 | -7.0965
3.9617 | -34.0468 | -7.1073
3.9678 | -34.047 | -7.118
3.9739 | -34.0473 | -7.1287
3.98 | -34.0475 | -7.1396
3.9861 | -34.0478 | -7.1504
3.9922 | -34.048 | -7.1611
3.9983 | -34.0483 | -7.1717
4.0044 | -34.0485 | -7.1825
4.0105 | -34.0488 | -7.1932
4.0166 | -34.049 | -7.204
4.0227 | -34.0493 | -7.2147
4.0288 | -34.0495 | -7.2256
4.0349 | -34.0497 | -7.2363
4.041 | -34.05 | -7.2469
4.0471 | -34.0502 | -7.2578
4.0532 | -34.0505 | -7.2685
4.0593 | -34.0508 | -7.2793
4.0654 | -34.051 | -7.29
4.0715 | -34.0512 | -7.3008
4.0776 | -34.0515 | -7.3113
4.0838 | -34.0518 | -7.3221
4.0899 | -34.052 | -7.3328
4.096 | -34.0522 | -7.3434
4.1021 | -34.0525 | -7.3541
4.1082 | -34.0527 | -7.3648
4.1143 | -34.053 | -7.3754
4.1204 | -34.0532 | -7.3861
4.1265 | -34.0535 | -7.3966
4.1326 | -34.0537 | -7.4072
4.1387 | -34.054 | -7.4178
4.1448 | -34.0542 | -7.4284
4.1509 | -34.0545 | -7.439
4.157 | -34.0547 | -7.4498
4.1631 | -34.055 | -7.4605
4.1692 | -34.0552 | -7.471
4.1753 | -34.0555 | -7.4817
4.1814 | -34.0557 | -7.4922
4.1875 | -34.056 | -7.5029
4.1936 | -34.0562 | -7.5136
4.1997 | -34.0564 | -7.5243
4.2058 | -34.0567 | -7.5349
4.2119 | -34.0569 | -7.5455
4.218 | -34.0572 | -7.556
4.2241 | -34.0574 | -7.5667
4.2303 | -34.0577 | -7.5772
4.2364 | -34.0579 | -7.5878
4.2425 | -34.0581 | -7.5984
4.2486 | -34.0584 | -7.6091
4.2547 | -34.0586 | -7.6196
4.2608 | -34.0589 | -7.6303
4.2669 | -34.0591 | -7.6409
4.273 | -34.0593 | -7.6515
4.2791 | -34.0596 | -7.6621
4.2852 | -34.0599 | -7.6726
4.2913 | -34.0601 | -7.6832
4.2974 | -34.0603 | -7.6937
4.3035 | -34.0606 | -7.7044
4.3096 | -34.0608 | -7.7151
4.3157 | -34.0611 | -7.7256
4.3218 | -34.0613 | -7.7361
4.3279 | -34.0616 | -7.7467
4.334 | -34.0618 | -7.7573
4.3401 | -34.062 | -7.7678
4.3462 | -34.0623 | -7.7784
4.3523 | -34.0625 | -7.7888
4.3584 | -34.0627 | -7.7992
4.3645 | -34.063 | -7.8098
4.3707 | -34.0632 | -7.8203
4.3768 | -34.0634 | -7.8309
4.3829 | -34.0637 | -7.8416
4.389 | -34.0639 | -7.8522
4.3951 | -34.0641 | -7.8627
4.4012 | -34.0644 | -7.8733
4.4073 | -34.0646 | -7.8838
4.4134 | -34.0649 | -7.8943
4.4195 | -34.0651 | -7.9047
4.4256 | -34.0653 | -7.9152
4.4317 | -34.0656 | -7.9257
4.4378 | -34.0658 | -7.9361
4.4439 | -34.066 | -7.9466
4.45 | -34.0662 | -7.9571
4.4561 | -34.0665 | -7.9677
4.4622 | -34.0667 | -7.9783
4.4683 | -34.0669 | -7.9888
4.4744 | -34.0672 | -7.9993
4.4805 | -34.0674 | -8.0097
4.4866 | -34.0676 | -8.0202
4.4927 | -34.0679 | -8.0308
4.4988 | -34.0681 | -8.0412
4.5049 | -34.0683 | -8.0516
4.511 | -34.0685 | -8.0621
4.5172 | -34.0688 | -8.0725
4.5233 | -34.069 | -8.0832
4.5294 | -34.0692 | -8.0936
4.5355 | -34.0694 | -8.1041
4.5416 | -34.0697 | -8.1145
4.5477 | -34.0699 | -8.125
4.5538 | -34.0701 | -8.1355
4.5599 | -34.0703 | -8.1459
4.566 | -34.0705 | -8.1564
4.5721 | -34.0708 | -8.1668
4.5782 | -34.071 | -8.1773
4.5843 | -34.0713 | -8.1879
4.5904 | -34.0715 | -8.1983
4.5965 | -34.0717 | -8.2088
4.6026 | -34.072 | -8.2193
4.6087 | -34.0722 | -8.2298
4.6148 | -34.0724 | -8.2404
4.6209 | -34.0727 | -8.2508
4.627 | -34.0729 | -8.2614
4.6331 | -34.0731 | -8.2719
4.6392 | -34.0733 | -8.2823
4.6453 | -34.0736 | -8.2927
4.6514 | -34.0738 | -8.3032
4.6576 | -34.074 | -8.3138
4.6637 | -34.0742 | -8.3243
4.6698 | -34.0745 | -8.3348
4.6759 | -34.0747 | -8.3453
4.682 | -34.0749 | -8.3557
4.6881 | -34.0752 | -8.3662
4.6942 | -34.0754 | -8.3767
4.7003 | -34.0756 | -8.3872
4.7064 | -34.0759 | -8.3975
4.7125 | -34.0761 | -8.408
4.7186 | -34.0763 | -8.4183
4.7247 | -34.0766 | -8.4288
4.7308 | -34.0768 | -8.4392
4.7369 | -34.077 | -8.4497
4.743 | -34.0773 | -8.4601
4.7491 | -34.0775 | -8.4706
4.7552 | -34.0778 | -8.4811
4.7613 | -34.078 | -8.4916
4.7674 | -34.0782 | -8.5021
4.7735 | -34.0784 | -8.5126
4.7796 | -34.0787 | -8.5231
4.7857 | -34.0789 | -8.5336
4.7918 | -34.0792 | -8.544
4.7979 | -34.0794 | -8.5544
4.8041 | -34.0796 | -8.5649
4.8102 | -34.0798 | -8.5753
4.8163 | -34.0801 | -8.5857
4.8224 | -34.0803 | -8.5962
4.8285 | -34.0806 | -8.6066
4.8346 | -34.0808 | -8.617
4.8407 | -34.081 | -8.6274
4.8468 | -34.0813 | -8.6379
4.8529 | -34.0815 | -8.6484
4.859 | -34.0818 | -8.6589
4.8651 | -34.082 | -8.6693
4.8712 | -34.0823 | -8.6797
4.8773 | -34.0825 | -8.6902
4.8834 | -34.0828 | -8.7007
4.8895 | -34.083 | -8.711
4.8956 | -34.0832 | -8.7215
4.9017 | -34.0835 | -8.7319
4.9078 | -34.0837 | -8.7424
4.9139 | -34.084 | -8.7527
4.92 | -34.0842 | -8.763
4.9261 | -34.0844 | -8.7734
4.9322 | -34.0847 | -8.7839
4.9383 | -34.0849 | -8.7943
4.9445 | -34.0852 | -8.8047
4.9506 | -34.0854 | -8.8151
4.9567 | -34.0857 | -8.8255
4.9628 | -34.0859 | -8.8359
4.9689 | -34.0862 | -8.8463
4.975 | -34.0864 | -8.8568
4.9811 | -34.0867 | -8.8672
4.9872 | -34.0869 | -8.8775
4.9933 | -34.0872 | -8.8879
4.9994 | -34.0874 | -8.8984
5.0055 | -34.0877 | -8.9088
5.0116 | -34.0879 | -8.9193
5.0177 | -34.0882 | -8.9297
5.0238 | -34.0884 | -8.9402
5.0299 | -34.0887 | -8.9506
5.036 | -34.089 | -8.9611
5.0421 | -34.0892 | -8.9715
5.0482 | -34.0895 | -8.9819
5.0543 | -34.0897 | -8.9923
5.0604 | -34.09 | -9.0026
5.0665 | -34.0902 | -9.013
5.0726 | -34.0905 | -9.0234
5.0787 | -34.0908 | -9.0339
5.0848 | -34.091 | -9.0442
5.091 | -34.0913 | -9.0546
5.0971 | -34.0915 | -9.0649
5.1032 | -34.0918 | -9.0754
5.1093 | -34.092 | -9.0857
5.1154 | -34.0923 | -9.0961
5.1215 | -34.0925 | -9.1065
5.1276 | -34.0928 | -9.1168
5.1337 | -34.0931 | -9.127
5.1398 | -34.0933 | -9.1374
5.1459 | -34.0936 | -9.1478
5.152 | -34.0938 | -9.1582
5.1581 | -34.0941 | -9.1687
5.1642 | -34.0943 | -9.1791
5.1703 | -34.0946 | -9.1895
5.1764 | -34.0949 | -9.1999
5.1825 | -34.0951 | -9.2102
5.1886 | -34.0954 | -9.2206
5.1947 | -34.0956 | -9.2311
5.2008 | -34.0959 | -9.2414
5.2069 | -34.0962 | -9.2519
5.213 | -34.0964 | -9.2623
5.2191 | -34.0967 | -9.2727
5.2252 | -34.097 | -9.2831
5.2314 | -34.0972 | -9.2935
5.2375 | -34.0975 | -9.3039
5.2436 | -34.0978 | -9.3144
5.2497 | -34.098 | -9.3247
5.2558 | -34.0983 | -9.3351
5.2619 | -34.0986 | -9.3456
5.268 | -34.0989 | -9.3559
5.2741 | -34.0991 | -9.3663
5.2802 | -34.0994 | -9.3768
5.2863 | -34.0997 | -9.3871
5.2924 | -34.1 | -9.3976
5.2985 | -34.1002 | -9.408
5.3046 | -34.1005 | -9.4184
5.3107 | -34.1008 | -9.4289
5.3168 | -34.1011 | -9.4393
5.3229 | -34.1013 | -9.4497
5.329 | -34.1016 | -9.4601
5.3351 | -34.1019 | -9.4704
5.3412 | -34.1022 | -9.4809
5.3473 | -34.1024 | -9.4913
5.3534 | -34.1027 | -9.5017
5.3595 | -34.103 | -9.5122
5.3656 | -34.1033 | -9.5226
5.3717 | -34.1036 | -9.5331
5.3779 | -34.1039 | -9.5435
5.384 | -34.1041 | -9.5539
5.3901 | -34.1044 | -9.5643
5.3962 | -34.1047 | -9.5747
5.4023 | -34.105 | -9.5851
5.4084 | -34.1052 | -9.5955
5.4145 | -34.1055 | -9.6059
5.4206 | -34.1058 | -9.6163
5.4267 | -34.1061 | -9.6267
5.4328 | -34.1064 | -9.6372
5.4389 | -34.1066 | -9.6476
5.445 | -34.1069 | -9.6581
5.4511 | -34.1072 | -9.6684
5.4572 | -34.1075 | -9.6788
5.4633 | -34.1078 | -9.6893
5.4694 | -34.1081 | -9.6997
5.4755 | -34.1083 | -9.7101
5.4816 | -34.1086 | -9.7205
5.4877 | -34.1089 | -9.7308
5.4938 | -34.1092 | -9.7412
5.4999 | -34.1095 | -9.7516
5.506 | -34.1098 | -9.762
5.5121 | -34.1101 | -9.7723
5.5183 | -34.1104 | -9.7828
5.5244 | -34.1107 | -9.7933
5.5305 | -34.1109 | -9.8037
5.5366 | -34.1112 | -9.8141
5.5427 | -34.1115 | -9.8246
5.5488 | -34.1118 | -9.835
5.5549 | -34.1121 | -9.8455
5.561 | -34.1124 | -9.856
5.5671 | -34.1127 | -9.8664
5.5732 | -34.113 | -9.8769
5.5793 | -34.1133 | -9.8875
5.5854 | -34.1136 | -9.8979
5.5915 | -34.1139 | -9.9083
5.5976 | -34.1142 | -9.9188
5.6037 | -34.1144 | -9.9293
5.6098 | -34.1147 | -9.9397
5.6159 | -34.115 | -9.9501
5.622 | -34.1153 | -9.9606
5.6281 | -34.1156 | -9.971
5.6342 | -34.1159 | -9.9815
5.6403 | -34.1162 | -9.9919
5.6464 | -34.1165 | -10.0024
5.6525 | -34.1168 | -10.0129
5.6586 | -34.1171 | -10.0232
5.6648 | -34.1174 | -10.0336
5.6709 | -34.1177 | -10.044
5.677 | -34.118 | -10.0545
5.6831 | -34.1183 | -10.0649
5.6892 | -34.1186 | -10.0755
5.6953 | -34.1189 | -10.0859
5.7014 | -34.1192 | -10.0964
5.7075 | -34.1195 | -10.1068
5.7136 | -34.1198 | -10.1173
5.7197 | -34.1201 | -10.1277
5.7258 | -34.1204 | -10.1382
5.7319 | -34.1207 | -10.1486
5.738 | -34.121 | -10.159
5.7441 | -34.1214 | -10.1694
5.7502 | -34.1217 | -10.1797
5.7563 | -34.1219 | -10.1902
5.7624 | -34.1222 | -10.2006
5.7685 | -34.1226 | -10.2111
5.7746 | -34.1228 | -10.2216
5.7807 | -34.1231 | -10.2321
5.7868 | -34.1235 | -10.2425
5.7929 | -34.1238 | -10.2529
5.799 | -34.1241 | -10.2634
5.8052 | -34.1244 | -10.2738
5.8113 | -34.1246 | -10.2842
5.8174 | -34.125 | -10.2947
5.8235 | -34.1253 | -10.305
5.8296 | -34.1256 | -10.3155
5.8357 | -34.1259 | -10.326
5.8418 | -34.1262 | -10.3365
5.8479 | -34.1265 | -10.347
5.854 | -34.1268 | -10.3574
5.8601 | -34.1271 | -10.3678
5.8662 | -34.1274 | -10.3782
5.8723 | -34.1277 | -10.3887
5.8784 | -34.128 | -10.3991
5.8845 | -34.1283 | -10.4095
5.8906 | -34.1286 | -10.4199
5.8967 | -34.1289 | -10.4303
5.9028 | -34.1292 | -10.4406
5.9089 | -34.1295 | -10.4509
5.915 | -34.1298 | -10.4613
5.9211 | -34.1301 | -10.4718
5.9272 | -34.1304 | -10.4821
5.9333 | -34.1307 | -10.4926
5.9394 | -34.131 | -10.5029
5.9455 | -34.1313 | -10.5132
5.9517 | -34.1316 | -10.5235
5.9578 | -34.1319 | -10.534
5.9639 | -34.1322 | -10.5442
5.97 | -34.1325 | -10.5545
5.9761 | -34.1328 | -10.5648
5.9822 | -34.1331 | -10.5752
5.9883 | -34.1334 | -10.5855
5.9944 | -34.1337 | -10.5959
6.0005 | -34.134 | -10.6062
6.0066 | -34.1343 | -10.6166
6.0127 | -34.1346 | -10.6268
6.0188 | -34.1349 | -10.6373
6.0249 | -34.1352 | -10.6477
6.031 | -34.1356 | -10.6582
6.0371 | -34.1359 | -10.6685
6.0432 | -34.1361 | -10.6789
6.0493 | -34.1365 | -10.689
6.0554 | -34.1368 | -10.6993
6.0615 | -34.1371 | -10.7096
6.0676 | -34.1374 | -10.72
6.0737 | -34.1377 | -10.7303
6.0798 | -34.138 | -10.7406
6.0859 | -34.1383 | -10.7508
6.0921 | -34.1386 | -10.7611
6.0982 | -34.1389 | -10.7715
6.1043 | -34.1392 | -10.7817
6.1104 | -34.1396 | -10.792
6.1165 | -34.1399 | -10.8021
6.1226 | -34.1402 | -10.8124
6.1287 | -34.1405 | -10.8227
6.1348 | -34.1408 | -10.8329
6.1409 | -34.1411 | -10.8433
6.147 | -34.1414 | -10.8535
6.1531 | -34.1417 | -10.8637
6.1592 | -34.142 | -10.8739
6.1653 | -34.1423 | -10.8841
6.1714 | -34.1426 | -10.8944
6.1775 | -34.1429 | -10.9046
6.1836 | -34.1432 | -10.9149
6.1897 | -34.1436 | -10.925
6.1958 | -34.1439 | -10.9353
6.2019 | -34.1442 | -10.9456
6.208 | -34.1445 | -10.9558
6.2141 | -34.1448 | -10.966
6.2202 | -34.1451 | -10.9763
6.2263 | -34.1454 | -10.9864
6.2325 | -34.1457 | -10.9965
6.2386 | -34.146 | -11.0066
6.2447 | -34.1463 | -11.0169
6.2508 | -34.1467 | -11.0271
6.2569 | -34.147 | -11.0372
6.263 | -34.1473 | -11.0475
6.2691 | -34.1476 | -11.0577
6.2752 | -34.1479 | -11.068
6.2813 | -34.1482 | -11.0782
6.2874 | -34.1485 | -11.0884
6.2935 | -34.1489 | -11.0985
6.2996 | -34.1492 | -11.1088
6.3057 | -34.1495 | -11.1189
6.3118 | -34.1498 | -11.1292
6.3179 | -34.1501 | -11.1394
6.324 | -34.1504 | -11.1495
6.3301 | -34.1507 | -11.1596
6.3362 | -34.151 | -11.1697
6.3423 | -34.1513 | -11.1799
6.3484 | -34.1517 | -11.1901
6.3545 | -34.152 | -11.2002
6.3606 | -34.1523 | -11.2103
6.3667 | -34.1526 | -11.2203
6.3728 | -34.1529 | -11.2305
6.379 | -34.1532 | -11.2407
6.3851 | -34.1535 | -11.2508
6.3912 | -34.1538 | -11.261
6.3973 | -34.1542 | -11.2711
6.4034 | -34.1545 | -11.2812
6.4095 | -34.1548 | -11.2914
6.4156 | -34.1551 | -11.3015
6.4217 | -34.1554 | -11.3117
6.4278 | -34.1557 | -11.3218
6.4339 | -34.1561 | -11.3319
6.44 | -34.1564 | -11.3419
6.4461 | -34.1567 | -11.3522
6.4522 | -34.157 | -11.3622
6.4583 | -34.1573 | -11.3725
6.4644 | -34.1576 | -11.3826
6.4705 | -34.1579 | -11.3926
6.4766 | -34.1583 | -11.4027
6.4827 | -34.1586 | -11.4129
6.4888 | -34.1589 | -11.423
6.4949 | -34.1592 | -11.433
6.501 | -34.1595 | -11.4431
6.5071 | -34.1598 | -11.453
6.5132 | -34.1602 | -11.4632
6.5194 | -34.1605 | -11.4732
6.5255 | -34.1608 | -11.4833
6.5316 | -34.1611 | -11.4934
6.5377 | -34.1614 | -11.5035
6.5438 | -34.1617 | -11.5135
6.5499 | -34.162 | -11.5235
6.556 | -34.1624 | -11.5336
6.5621 | -34.1627 | -11.5438
6.5682 | -34.163 | -11.5538
6.5743 | -34.1633 | -11.5638
6.5804 | -34.1637 | -11.5739
6.5865 | -34.164 | -11.5839
6.5926 | -34.1643 | -11.594
6.5987 | -34.1646 | -11.604
6.6048 | -34.1649 | -11.614
6.6109 | -34.1652 | -11.6239
6.617 | -34.1656 | -11.634
6.6231 | -34.1659 | -11.6441
6.6292 | -34.1662 | -11.6541
6.6353 | -34.1665 | -11.6642
6.6414 | -34.1669 | -11.6742
6.6475 | -34.1672 | -11.6841
6.6536 | -34.1675 | -11.6941
6.6597 | -34.1678 | -11.7041
6.6659 | -34.1681 | -11.7141
6.672 | -34.1685 | -11.7241
6.6781 | -34.1688 | -11.7341
6.6842 | -34.1691 | -11.7441
6.6903 | -34.1694 | -11.7541
6.6964 | -34.1698 | -11.7642
6.7025 | -34.1701 | -11.7742
6.7086 | -34.1704 | -11.7842
6.7147 | -34.1707 | -11.7943
6.7208 | -34.171 | -11.8043
6.7269 | -34.1714 | -11.8143
6.733 | -34.1717 | -11.8243
6.7391 | -34.172 | -11.8343
6.7452 | -34.1723 | -11.8443
6.7513 | -34.1726 | -11.8542
6.7574 | -34.173 | -11.8642
6.7635 | -34.1733 | -11.8742
6.7696 | -34.1736 | -11.8842
6.7757 | -34.1739 | -11.8942
6.7818 | -34.1743 | -11.9041
6.7879 | -34.1746 | -11.914
6.794 | -34.1749 | -11.924
6.8001 | -34.1752 | -11.9339
6.8063 | -34.1755 | -11.9439
6.8124 | -34.1759 | -11.9539
6.8185 | -34.1762 | -11.9639
6.8246 | -34.1765 | -11.9738
6.8307 | -34.1769 | -11.9838
6.8368 | -34.1772 | -11.9938
6.8429 | -34.1775 | -12.0038
6.849 | -34.1779 | -12.0138
6.8551 | -34.1782 | -12.0237
6.8612 | -34.1785 | -12.0336
6.8673 | -34.1788 | -12.0436
6.8734 | -34.1791 | -12.0536
6.8795 | -34.1795 | -12.0636
6.8856 | -34.1798 | -12.0736
6.8917 | -34.1801 | -12.0835
6.8978 | -34.1805 | -12.0934
6.9039 | -34.1808 | -12.1034
6.91 | -34.1811 | -12.1133
6.9161 | -34.1814 | -12.1232
6.9222 | -34.1817 | -12.1332
6.9283 | -34.1821 | -12.143
6.9344 | -34.1824 | -12.1531
6.9405 | -34.1827 | -12.163
6.9466 | -34.1831 | -12.1729
6.9528 | -34.1834 | -12.1829
6.9589 | -34.1837 | -12.1929
6.965 | -34.1841 | -12.2029
6.9711 | -34.1844 | -12.2128
6.9772 | -34.1847 | -12.2227
6.9833 | -34.1851 | -12.2327
6.9894 | -34.1854 | -12.2428
6.9955 | -34.1857 | -12.2526
7.0016 | -34.186 | -12.2625
7.0077 | -34.1864 | -12.2725
7.0138 | -34.1867 | -12.2824
7.0199 | -34.187 | -12.2924
7.026 | -34.1874 | -12.3023
7.0321 | -34.1877 | -12.3122
7.0382 | -34.188 | -12.322
7.0443 | -34.1884 | -12.332
7.0504 | -34.1887 | -12.3421
7.0565 | -34.189 | -12.3521
7.0626 | -34.1893 | -12.362
7.0687 | -34.1897 | -12.3718
7.0748 | -34.19 | -12.3817
7.0809 | -34.1903 | -12.3917
7.087 | -34.1907 | -12.4016
7.0932 | -34.191 | -12.4116
7.0993 | -34.1913 | -12.4215
7.1054 | -34.1917 | -12.4314
7.1115 | -34.192 | -12.4413
7.1176 | -34.1923 | -12.4512
7.1237 | -34.1927 | -12.4612
7.1298 | -34.193 | -12.4711
7.1359 | -34.1933 | -12.481
7.142 | -34.1937 | -12.4909
7.1481 | -34.194 | -12.5009
7.1542 | -34.1943 | -12.5109
7.1603 | -34.1947 | -12.5208
7.1664 | -34.195 | -12.5307
7.1725 | -34.1953 | -12.5407
7.1786 | -34.1957 | -12.5506
7.1847 | -34.196 | -12.5606
7.1908 | -34.1963 | -12.5706
7.1969 | -34.1967 | -12.5806
7.203 | -34.197 | -12.5905
7.2091 | -34.1973 | -12.6005
7.2152 | -34.1977 | -12.6105
7.2213 | -34.198 | -12.6205
7.2274 | -34.1983 | -12.6303
7.2335 | -34.1986 | -12.6403
7.2397 | -34.199 | -12.6503
7.2458 | -34.1993 | -12.6602
7.2519 | -34.1996 | -12.6701
7.258 | -34.2 | -12.6801
7.2641 | -34.2003 | -12.69
7.2702 | -34.2006 | -12.7
7.2763 | -34.201 | -12.7099
7.2824 | -34.2013 | -12.7199
7.2885 | -34.2017 | -12.7299
7.2946 | -34.202 | -12.7398
7.3007 | -34.2023 | -12.7499
7.3068 | -34.2027 | -12.7599
7.3129 | -34.203 | -12.7699
7.319 | -34.2033 | -12.7799
7.3251 | -34.2037 | -12.7898
7.3312 | -34.204 | -12.7997
7.3373 | -34.2043 | -12.8096
7.3434 | -34.2047 | -12.8195
7.3495 | -34.205 | -12.8295
7.3556 | -34.2054 | -12.8395
7.3617 | -34.2057 | -12.8494
7.3678 | -34.206 | -12.8594
7.3739 | -34.2064 | -12.8695
7.3801 | -34.2067 | -12.8795
7.3862 | -34.207 | -12.8894
7.3923 | -34.2074 | -12.8994
7.3984 | -34.2077 | -12.9093
7.4045 | -34.2081 | -12.9192
7.4106 | -34.2084 | -12.9291
7.4167 | -34.2088 | -12.939
7.4228 | -34.2091 | -12.9489
7.4289 | -34.2095 | -12.9588
7.435 | -34.2098 | -12.9687
7.4411 | -34.2101 | -12.9788
7.4472 | -34.2104 | -12.9887
7.4533 | -34.2108 | -12.9987
7.4594 | -34.2111 | -13.0087
7.4655 | -34.2115 | -13.0186
7.4716 | -34.2118 | -13.0285
7.4777 | -34.2121 | -13.0385
7.4838 | -34.2125 | -13.0485
7.4899 | -34.2128 | -13.0585
7.496 | -34.2132 | -13.0684
7.5021 | -34.2135 | -13.0783
7.5082 | -34.2138 | -13.0883
7.5143 | -34.2142 | -13.0982
7.5204 | -34.2145 | -13.1082
7.5266 | -34.2149 | -13.1182
7.5327 | -34.2152 | -13.1281
7.5388 | -34.2155 | -13.1382
7.5449 | -34.2159 | -13.1481
7.551 | -34.2162 | -13.1581
7.5571 | -34.2166 | -13.168
7.5632 | -34.217 | -13.178
7.5693 | -34.2173 | -13.1879
7.5754 | -34.2176 | -13.1979
7.5815 | -34.218 | -13.2078
7.5876 | -34.2183 | -13.2178
7.5937 | -34.2187 | -13.2277
7.5998 | -34.219 | -13.2376
7.6059 | -34.2194 | -13.2475
7.612 | -34.2197 | -13.2574
7.6181 | -34.2201 | -13.2673
7.6242 | -34.2204 | -13.2773
7.6303 | -34.2208 | -13.2872
7.6364 | -34.2212 | -13.2972
7.6425 | -34.2215 | -13.3071
7.6486 | -34.2218 | -13.3171
7.6547 | -34.2222 | -13.327
7.6608 | -34.2225 | -13.337
7.667 | -34.2229 | -13.347
7.6731 | -34.2232 | -13.357
7.6792 | -34.2236 | -13.3669
7.6853 | -34.224 | -13.3769
7.6914 | -34.2243 | -13.3869
7.6975 | -34.2247 | -13.3968
7.7036 | -34.225 | -13.4067
7.7097 | -34.2254 | -13.4166
7.7158 | -34.2257 | -13.4266
7.7219 | -34.2261 | -13.4366
7.728 | -34.2264 | -13.4466
7.7341 | -34.2268 | -13.4566
7.7402 | -34.2271 | -13.4665
7.7463 | -34.2275 | -13.4764
7.7524 | -34.2278 | -13.4864
7.7585 | -34.2282 | -13.4964
7.7646 | -34.2285 | -13.5063
7.7707 | -34.2289 | -13.5163
7.7768 | -34.2293 | -13.5263
7.7829 | -34.2296 | -13.5363
7.789 | -34.23 | -13.5464
7.7951 | -34.2303 | -13.5564
7.8012 | -34.2307 | -13.5663
7.8073 | -34.2311 | -13.5762
7.8135 | -34.2314 | -13.5861
7.8196 | -34.2318 | -13.5961
7.8257 | -34.2321 | -13.606
7.8318 | -34.2325 | -13.616
7.8379 | -34.2328 | -13.6259
7.844 | -34.2332 | -13.6356
7.8501 | -34.2336 | -13.6456
7.8562 | -34.2339 | -13.6555
7.8623 | -34.2343 | -13.6654
7.8684 | -34.2347 | -13.6755
7.8745 | -34.2351 | -13.6854
7.8806 | -34.2354 | -13.6953
7.8867 | -34.2358 | -13.7053
7.8928 | -34.2362 | -13.7151
7.8989 | -34.2365 | -13.725
7.905 | -34.2369 | -13.7349
7.9111 | -34.2373 | -13.7449
7.9172 | -34.2376 | -13.7547
7.9233 | -34.238 | -13.7647
7.9294 | -34.2384 | -13.7746
7.9355 | -34.2387 | -13.7845
7.9416 | -34.2391 | -13.7945
7.9477 | -34.2395 | -13.8045
7.9539 | -34.2398 | -13.8145
7.96 | -34.2402 | -13.8244
7.9661 | -34.2406 | -13.8344
7.9722 | -34.241 | -13.8442
7.9783 | -34.2413 | -13.8542
7.9844 | -34.2418 | -13.8641
7.9905 | -34.2421 | -13.8741
7.9966 | -34.2425 | -13.884
8.0027 | -34.2429 | -13.894
8.0088 | -34.2432 | -13.9038
8.0149 | -34.2436 | -13.9137
8.021 | -34.2439 | -13.9235
8.0271 | -34.2443 | -13.9336
8.0332 | -34.2447 | -13.9435
8.0393 | -34.245 | -13.9534
8.0454 | -34.2454 | -13.9634
8.0515 | -34.2458 | -13.9731
8.0576 | -34.2461 | -13.9831
8.0637 | -34.2465 | -13.9931
8.0698 | -34.2469 | -14.0031
8.0759 | -34.2472 | -14.0129
8.082 | -34.2476 | -14.0227
8.0881 | -34.2479 | -14.0327
8.0942 | -34.2483 | -14.0425
8.1004 | -34.2487 | -14.0524
8.1065 | -34.2491 | -14.0622
8.1126 | -34.2494 | -14.0722
8.1187 | -34.2498 | -14.0821
8.1248 | -34.2502 | -14.0918
8.1309 | -34.2506 | -14.1017
8.137 | -34.2509 | -14.1115
8.1431 | -34.2513 | -14.1214
8.1492 | -34.2517 | -14.1313
8.1553 | -34.2521 | -14.1412
8.1614 | -34.2524 | -14.151
8.1675 | -34.2528 | -14.161
8.1736 | -34.2532 | -14.171
8.1797 | -34.2536 | -14.181
8.1858 | -34.2539 | -14.1908
8.1919 | -34.2543 | -14.2005
8.198 | -34.2547 | -14.2103
8.2041 | -34.2551 | -14.2201
8.2102 | -34.2554 | -14.2299
8.2163 | -34.2558 | -14.2397
8.2224 | -34.2562 | -14.2496
8.2285 | -34.2566 | -14.2593
8.2346 | -34.2569 | -14.2693
8.2408 | -34.2573 | -14.2791
8.2469 | -34.2577 | -14.289
8.253 | -34.2581 | -14.2988
8.2591 | -34.2585 | -14.3088
8.2652 | -34.2589 | -14.3184
8.2713 | -34.2593 | -14.3283
8.2774 | -34.2596 | -14.3382
8.2835 | -34.26 | -14.3481
8.2896 | -34.2604 | -14.3578
8.2957 | -34.2608 | -14.3677
8.3018 | -34.2612 | -14.3775
8.3079 | -34.2616 | -14.3872
8.314 | -34.2619 | -14.3969
8.3201 | -34.2623 | -14.4068
8.3262 | -34.2627 | -14.4166
8.3323 | -34.2631 | -14.4263
8.3384 | -34.2635 | -14.4361
8.3445 | -34.2639 | -14.446
8.3506 | -34.2642 | -14.4559
8.3567 | -34.2646 | -14.4657
8.3628 | -34.265 | -14.4756
8.3689 | -34.2654 | -14.4852
8.375 | -34.2658 | -14.4952
8.3812 | -34.2661 | -14.5051
8.3873 | -34.2665 | -14.5149
8.3934 | -34.2669 | -14.5248
8.3995 | -34.2673 | -14.5346
8.4056 | -34.2677 | -14.5444
8.4117 | -34.2681 | -14.5542
8.4178 | -34.2685 | -14.5641
8.4239 | -34.2688 | -14.5739
8.43 | -34.2692 | -14.5837
8.4361 | -34.2696 | -14.5936
8.4422 | -34.27 | -14.6032
8.4483 | -34.2704 | -14.6129
8.4544 | -34.2708 | -14.6227
8.4605 | -34.2711 | -14.6325
8.4666 | -34.2715 | -14.6423
8.4727 | -34.2719 | -14.6521
8.4788 | -34.2723 | -14.6618
8.4849 | -34.2727 | -14.6717
8.491 | -34.2731 | -14.6816
8.4971 | -34.2735 | -14.6913
8.5032 | -34.2739 | -14.7011
8.5093 | -34.2742 | -14.7108
8.5154 | -34.2746 | -14.7205
8.5215 | -34.275 | -14.7302
8.5277 | -34.2754 | -14.7401
8.5338 | -34.2758 | -14.7499
8.5399 | -34.2762 | -14.7597
8.546 | -34.2766 | -14.7693
8.5521 | -34.2769 | -14.7791
8.5582 | -34.2773 | -14.7889
8.5643 | -34.2777 | -14.7985
8.5704 | -34.2781 | -14.8082
8.5765 | -34.2785 | -14.8179
8.5826 | -34.2789 | -14.8275
8.5887 | -34.2792 | -14.8373
8.5948 | -34.2796 | -14.847
8.6009 | -34.28 | -14.8568
8.607 | -34.2804 | -14.8667
8.6131 | -34.2808 | -14.8764
8.6192 | -34.2812 | -14.8859
8.6253 | -34.2815 | -14.8957
8.6314 | -34.2819 | -14.9055
8.6375 | -34.2823 | -14.9153
8.6436 | -34.2827 | -14.925
8.6497 | -34.2831 | -14.9345
8.6558 | -34.2835 | -14.9443
8.6619 | -34.2838 | -14.954
8.6681 | -34.2842 | -14.9636
8.6742 | -34.2846 | -14.9733
8.6803 | -34.285 | -14.9832
8.6864 | -34.2854 | -14.9927
8.6925 | -34.2858 | -15.0025
8.6986 | -34.2861 | -15.0122
8.7047 | -34.2865 | -15.0221
8.7108 | -34.2869 | -15.0317
8.7169 | -34.2873 | -15.0414
8.723 | -34.2877 | -15.051
8.7291 | -34.2881 | -15.0607
8.7352 | -34.2884 | -15.0706
8.7413 | -34.2888 | -15.0802
8.7474 | -34.2892 | -15.0899
8.7535 | -34.2896 | -15.0996
8.7596 | -34.29 | -15.1092
8.7657 | -34.2904 | -15.119
8.7718 | -34.2908 | -15.1287
8.7779 | -34.2912 | -15.1384
8.784 | -34.2916 | -15.1481
8.7901 | -34.2919 | -15.1577
8.7962 | -34.2923 | -15.1674
8.8023 | -34.2927 | -15.1771
8.8084 | -34.2931 | -15.1868
8.8146 | -34.2935 | -15.1965
8.8207 | -34.2939 | -15.2061
8.8268 | -34.2943 | -15.2156
8.8329 | -34.2946 | -15.2253
8.839 | -34.295 | -15.2349
8.8451 | -34.2954 | -15.2446
8.8512 | -34.2958 | -15.2543
8.8573 | -34.2962 | -15.264
8.8634 | -34.2966 | -15.2736
8.8695 | -34.2969 | -15.2833
8.8756 | -34.2973 | -15.2931
8.8817 | -34.2977 | -15.3027
8.8878 | -34.2981 | -15.3125
8.8939 | -34.2985 | -15.3221
8.9 | -34.2989 | -15.3317
8.9061 | -34.2993 | -15.3414
8.9122 | -34.2996 | -15.351
8.9183 | -34.3 | -15.3606
8.9244 | -34.3004 | -15.3702
8.9305 | -34.3008 | -15.3797
8.9366 | -34.3012 | -15.3894
8.9427 | -34.3016 | -15.3991
8.9488 | -34.302 | -15.4087
8.955 | -34.3024 | -15.4184
8.9611 | -34.3028 | -15.428
8.9672 | -34.3032 | -15.4375
8.9733 | -34.3036 | -15.4472
8.9794 | -34.3039 | -15.4567
8.9855 | -34.3043 | -15.4663
8.9916 | -34.3047 | -15.476
8.9977 | -34.3051 | -15.4858
9.0038 | -34.3055 | -15.4953
9.0099 | -34.3059 | -15.505
9.016 | -34.3063 | -15.5147
9.0221 | -34.3067 | -15.5243
9.0282 | -34.3071 | -15.5339
9.0343 | -34.3075 | -15.5436
9.0404 | -34.3079 | -15.5532
9.0465 | -34.3083 | -15.5628
9.0526 | -34.3086 | -15.5724
9.0587 | -34.309 | -15.5819
9.0648 | -34.3094 | -15.5916
9.0709 | -34.3098 | -15.6012
9.077 | -34.3102 | -15.6107
9.0831 | -34.3106 | -15.6205
9.0892 | -34.311 | -15.6301
9.0953 | -34.3114 | -15.6397
9.1015 | -34.3118 | -15.6494
9.1076 | -34.3122 | -15.6588
9.1137 | -34.3126 | -15.6684
9.1198 | -34.313 | -15.6781
9.1259 | -34.3134 | -15.6878
9.132 | -34.3138 | -15.6974
9.1381 | -34.3141 | -15.7069
9.1442 | -34.3145 | -15.7164
9.1503 | -34.3149 | -15.7261
9.1564 | -34.3153 | -15.7357
9.1625 | -34.3157 | -15.7453
9.1686 | -34.3161 | -15.7549
9.1747 | -34.3165 | -15.7645
9.1808 | -34.3169 | -15.774
9.1869 | -34.3173 | -15.7835
9.193 | -34.3177 | -15.7931
9.1991 | -34.3181 | -15.8027
9.2052 | -34.3185 | -15.8123
9.2113 | -34.3189 | -15.8218
9.2174 | -34.3193 | -15.8313
9.2235 | -34.3197 | -15.8409
9.2296 | -34.3201 | -15.8504
9.2357 | -34.3205 | -15.86
9.2419 | -34.3209 | -15.8696
9.248 | -34.3213 | -15.879
9.2541 | -34.3217 | -15.8886
9.2602 | -34.3221 | -15.8982
9.2663 | -34.3225 | -15.9077
9.2724 | -34.3229 | -15.9173
9.2785 | -34.3233 | -15.9269
9.2846 | -34.3237 | -15.9363
9.2907 | -34.3241 | -15.9459
9.2968 | -34.3245 | -15.9555
9.3029 | -34.3249 | -15.9652
9.309 | -34.3253 | -15.9748
9.3151 | -34.3257 | -15.9844
9.3212 | -34.3261 | -15.9939
9.3273 | -34.3265 | -16.0034
9.3334 | -34.3269 | -16.0129
9.3395 | -34.3273 | -16.0225
9.3456 | -34.3277 | -16.032
9.3517 | -34.3281 | -16.0416
9.3578 | -34.3285 | -16.0511
9.3639 | -34.3289 | -16.0606
9.37 | -34.3293 | -16.0702
9.3761 | -34.3297 | -16.0797
9.3822 | -34.3301 | -16.0892
9.3884 | -34.3305 | -16.0986
9.3945 | -34.3309 | -16.1082
9.4006 | -34.3313 | -16.1177
9.4067 | -34.3317 | -16.1274
9.4128 | -34.3321 | -16.137
9.4189 | -34.3325 | -16.1466
9.425 | -34.3329 | -16.156
9.4311 | -34.3333 | -16.1657
9.4372 | -34.3337 | -16.1752
9.4433 | -34.3341 | -16.1848
9.4494 | -34.3345 | -16.1943
9.4555 | -34.3349 | -16.2039
9.4616 | -34.3354 | -16.2134
9.4677 | -34.3358 | -16.223
9.4738 | -34.3362 | -16.2326
9.4799 | -34.3366 | -16.2421
9.486 | -34.337 | -16.2517
9.4921 | -34.3374 | -16.2613
9.4982 | -34.3378 | -16.2708
9.5043 | -34.3382 | -16.2803
9.5104 | -34.3386 | -16.2898
9.5165 | -34.339 | -16.2993
9.5226 | -34.3394 | -16.3089
9.5288 | -34.3398 | -16.3182
9.5349 | -34.3402 | -16.3277
9.541 | -34.3406 | -16.3372
9.5471 | -34.341 | -16.3468
9.5532 | -34.3414 | -16.3563
9.5593 | -34.3418 | -16.3658
9.5654 | -34.3423 | -16.3752
9.5715 | -34.3427 | -16.3847
9.5776 | -34.3431 | -16.3943
9.5837 | -34.3435 | -16.4039
9.5898 | -34.3439 | -16.4135
9.5959 | -34.3443 | -16.4229
9.602 | -34.3447 | -16.4324
9.6081 | -34.3451 | -16.4419
9.6142 | -34.3455 | -16.4514
9.6203 | -34.3459 | -16.4609
9.6264 | -34.3463 | -16.4705
9.6325 | -34.3467 | -16.48
9.6386 | -34.3471 | -16.4894
9.6447 | -34.3475 | -16.4989
9.6508 | -34.3479 | -16.5084
9.6569 | -34.3483 | -16.5179
9.663 | -34.3487 | -16.5274
9.6691 | -34.3491 | -16.5369
9.6753 | -34.3495 | -16.5464
9.6814 | -34.3499 | -16.5559
9.6875 | -34.3503 | -16.5654
9.6936 | -34.3507 | -16.5748
9.6997 | -34.3511 | -16.5844
9.7058 | -34.3516 | -16.5937
9.7119 | -34.352 | -16.6033
9.718 | -34.3524 | -16.6128
9.7241 | -34.3528 | -16.6223
9.7302 | -34.3532 | -16.6319
9.7363 | -34.3536 | -16.6413
9.7424 | -34.354 | -16.6507
9.7485 | -34.3544 | -16.6603
9.7546 | -34.3548 | -16.6698
9.7607 | -34.3552 | -16.6795
9.7668 | -34.3556 | -16.6889
9.7729 | -34.356 | -16.6985
9.779 | -34.3565 | -16.7078
9.7851 | -34.3569 | -16.7173
9.7912 | -34.3573 | -16.7268
9.7973 | -34.3577 | -16.7363
9.8034 | -34.3581 | -16.7458
9.8095 | -34.3585 | -16.7554
9.8157 | -34.359 | -16.7647
9.8218 | -34.3594 | -16.7743
9.8279 | -34.3598 | -16.7838
9.834 | -34.3602 | -16.7933
9.8401 | -34.3606 | -16.8027
9.8462 | -34.361 | -16.8121
9.8523 | -34.3614 | -16.8217
9.8584 | -34.3618 | -16.8312
9.8645 | -34.3622 | -16.8406
9.8706 | -34.3627 | -16.8501
9.8767 | -34.3631 | -16.8596
9.8828 | -34.3635 | -16.8691
9.8889 | -34.3639 | -16.8785
9.895 | -34.3643 | -16.888
9.9011 | -34.3647 | -16.8974
9.9072 | -34.3651 | -16.9068
9.9133 | -34.3655 | -16.9164
9.9194 | -34.366 | -16.9257
9.9255 | -34.3664 | -16.9351
9.9316 | -34.3668 | -16.9446
9.9377 | -34.3672 | -16.954
9.9438 | -34.3676 | -16.9636
9.9499 | -34.368 | -16.9731
9.956 | -34.3685 | -16.9825
9.9622 | -34.3689 | -16.9919
9.9683 | -34.3693 | -17.0013
9.9744 | -34.3697 | -17.0109
9.9805 | -34.3701 | -17.0203
9.9866 | -34.3706 | -17.0297
9.9927 | -34.371 | -17.0392
9.9988 | -34.3714 | -17.0486
10.0049 | -34.3718 | -17.058
10.011 | -34.3722 | -17.0676
10.0171 | -34.3726 | -17.077
10.0232 | -34.3731 | -17.0864
10.0293 | -34.3735 | -17.0958
10.0354 | -34.3739 | -17.1052
10.0415 | -34.3743 | -17.1146
10.0476 | -34.3747 | -17.1241
10.0537 | -34.3752 | -17.1336
10.0598 | -34.3756 | -17.143
10.0659 | -34.376 | -17.1525
10.072 | -34.3764 | -17.1619
10.0781 | -34.3768 | -17.1714
10.0842 | -34.3772 | -17.1808
10.0903 | -34.3777 | -17.1902
10.0964 | -34.3781 | -17.1995
10.1026 | -34.3785 | -17.209
10.1087 | -34.3789 | -17.2185
10.1148 | -34.3793 | -17.228
10.1209 | -34.3798 | -17.2374
10.127 | -34.3802 | -17.2467
10.1331 | -34.3806 | -17.2562
10.1392 | -34.381 | -17.2656
10.1453 | -34.3814 | -17.275
10.1514 | -34.3818 | -17.2844
10.1575 | -34.3823 | -17.2939
10.1636 | -34.3827 | -17.3032
10.1697 | -34.3831 | -17.3126
10.1758 | -34.3835 | -17.322
10.1819 | -34.3839 | -17.3314
10.188 | -34.3844 | -17.3408
10.1941 | -34.3848 | -17.3503
10.2002 | -34.3852 | -17.3597
10.2063 | -34.3856 | -17.3691
10.2124 | -34.3861 | -17.3785
10.2185 | -34.3865 | -17.3879
10.2246 | -34.3869 | -17.3973
10.2307 | -34.3873 | -17.4067
10.2368 | -34.3877 | -17.416
10.2429 | -34.3882 | -17.4254
10.2491 | -34.3886 | -17.4349
10.2552 | -34.389 | -17.4443
10.2613 | -34.3894 | -17.4537
10.2674 | -34.3899 | -17.4631
10.2735 | -34.3903 | -17.4725
10.2796 | -34.3907 | -17.4818
10.2857 | -34.3912 | -17.4912
10.2918 | -34.3916 | -17.5006
10.2979 | -34.392 | -17.51
10.304 | -34.3924 | -17.5191
10.3101 | -34.3929 | -17.5286
10.3162 | -34.3933 | -17.538
10.3223 | -34.3937 | -17.5473
10.3284 | -34.3941 | -17.5567
10.3345 | -34.3945 | -17.5661
10.3406 | -34.395 | -17.5753
10.3467 | -34.3954 | -17.5848
10.3528 | -34.3958 | -17.5942
10.3589 | -34.3962 | -17.6035
10.365 | -34.3967 | -17.6128
10.3711 | -34.3971 | -17.6223
10.3772 | -34.3975 | -17.6315
10.3833 | -34.3979 | -17.6408
10.3895 | -34.3984 | -17.6501
10.3956 | -34.3988 | -17.6595
10.4017 | -34.3992 | -17.6688
10.4078 | -34.3996 | -17.6782
10.4139 | -34.4001 | -17.6875
10.42 | -34.4005 | -17.6969
10.4261 | -34.4009 | -17.7063
10.4322 | -34.4013 | -17.7157
10.4383 | -34.4018 | -17.7251
10.4444 | -34.4022 | -17.7343
10.4505 | -34.4026 | -17.7437
10.4566 | -34.403 | -17.753
10.4627 | -34.4035 | -17.7623
10.4688 | -34.4039 | -17.7716
10.4749 | -34.4043 | -17.7809
10.481 | -34.4048 | -17.7903
10.4871 | -34.4052 | -17.7996
10.4932 | -34.4056 | -17.8089
10.4993 | -34.406 | -17.8182
10.5054 | -34.4065 | -17.8275
10.5115 | -34.4069 | -17.8369
10.5176 | -34.4073 | -17.8461
10.5237 | -34.4077 | -17.8555
10.5298 | -34.4081 | -17.8648
10.536 | -34.4086 | -17.8741
10.5421 | -34.409 | -17.8835
10.5482 | -34.4094 | -17.8929
10.5543 | -34.4099 | -17.9022
10.5604 | -34.4103 | -17.9115
10.5665 | -34.4107 | -17.9208
10.5726 | -34.4111 | -17.9301
10.5787 | -34.4115 | -17.9394
10.5848 | -34.412 | -17.9486
10.5909 | -34.4124 | -17.958
10.597 | -34.4128 | -17.9673
10.6031 | -34.4133 | -17.9767
10.6092 | -34.4137 | -17.986
10.6153 | -34.4141 | -17.9954
10.6214 | -34.4146 | -18.0046
10.6275 | -34.415 | -18.014
10.6336 | -34.4154 | -18.0232
10.6397 | -34.4158 | -18.0326
10.6458 | -34.4163 | -18.0418
10.6519 | -34.4167 | -18.0512
10.658 | -34.4171 | -18.0605
10.6641 | -34.4176 | -18.0697
10.6702 | -34.418 | -18.0791
10.6764 | -34.4184 | -18.0884
10.6825 | -34.4189 | -18.0976
10.6886 | -34.4193 | -18.1069
10.6947 | -34.4197 | -18.1161
10.7008 | -34.4201 | -18.1254
10.7069 | -34.4206 | -18.1346
10.713 | -34.421 | -18.1439
10.7191 | -34.4214 | -18.1531
10.7252 | -34.4218 | -18.1623
10.7313 | -34.4222 | -18.1717
10.7374 | -34.4227 | -18.181
10.7435 | -34.4231 | -18.1903
10.7496 | -34.4235 | -18.1996
10.7557 | -34.424 | -18.2088
10.7618 | -34.4244 | -18.2179
10.7679 | -34.4248 | -18.2272
10.774 | -34.4253 | -18.2365
10.7801 | -34.4257 | -18.2459
10.7862 | -34.4261 | -18.2551
10.7923 | -34.4265 | -18.2644
10.7984 | -34.427 | -18.2735
10.8045 | -34.4274 | -18.2827
10.8106 | -34.4278 | -18.292
10.8168 | -34.4283 | -18.3014
10.8229 | -34.4287 | -18.3107
10.829 | -34.4291 | -18.32
10.8351 | -34.4295 | -18.3291
10.8412 | -34.43 | -18.3383
10.8473 | -34.4304 | -18.3475
10.8534 | -34.4308 | -18.3568
10.8595 | -34.4312 | -18.366
10.8656 | -34.4317 | -18.3751
10.8717 | -34.4321 | -18.3844
10.8778 | -34.4326 | -18.3936
10.8839 | -34.433 | -18.4029
10.89 | -34.4334 | -18.4122
10.8961 | -34.4339 | -18.4213
10.9022 | -34.4343 | -18.4305
10.9083 | -34.4347 | -18.4397
10.9144 | -34.4352 | -18.449
10.9205 | -34.4356 | -18.4582
10.9266 | -34.436 | -18.4674
10.9327 | -34.4364 | -18.4766
10.9388 | -34.4369 | -18.4857
10.9449 | -34.4373 | -18.4949
10.951 | -34.4377 | -18.5041
10.9571 | -34.4382 | -18.5133
10.9633 | -34.4386 | -18.5226
10.9694 | -34.439 | -18.5319
10.9755 | -34.4395 | -18.541
10.9816 | -34.4399 | -18.5503
10.9877 | -34.4403 | -18.5594
10.9938 | -34.4408 | -18.5685
10.9999 | -34.4412 | -18.5777
11.006 | -34.4417 | -18.5868
11.0121 | -34.4421 | -18.596
11.0182 | -34.4426 | -18.6051
11.0243 | -34.443 | -18.6143
11.0304 | -34.4434 | -18.6235
11.0365 | -34.4438 | -18.6327
11.0426 | -34.4443 | -18.6417
11.0487 | -34.4447 | -18.651
11.0548 | -34.4451 | -18.6601
11.0609 | -34.4456 | -18.6693
11.067 | -34.446 | -18.6785
11.0731 | -34.4464 | -18.6876
11.0792 | -34.4469 | -18.6967
11.0853 | -34.4473 | -18.7058
11.0914 | -34.4478 | -18.715
11.0975 | -34.4482 | -18.7243
11.1037 | -34.4486 | -18.7334
11.1098 | -34.4491 | -18.7426
11.1159 | -34.4495 | -18.7517
11.122 | -34.4499 | -18.7609
11.1281 | -34.4503 | -18.77
11.1342 | -34.4508 | -18.7792
11.1403 | -34.4512 | -18.7884
11.1464 | -34.4517 | -18.7976
11.1525 | -34.4521 | -18.8066
11.1586 | -34.4526 | -18.8159
11.1647 | -34.453 | -18.8249
11.1708 | -34.4534 | -18.8341
11.1769 | -34.4539 | -18.8434
11.183 | -34.4543 | -18.8524
11.1891 | -34.4547 | -18.8614
11.1952 | -34.4552 | -18.8705
11.2013 | -34.4556 | -18.8797
11.2074 | -34.456 | -18.8889
11.2135 | -34.4565 | -18.8981
11.2196 | -34.4569 | -18.9072
11.2257 | -34.4574 | -18.9164
11.2318 | -34.4578 | -18.9255
11.2379 | -34.4583 | -18.9346
11.244 | -34.4587 | -18.9438
11.2502 | -34.4591 | -18.9529
11.2563 | -34.4596 | -18.9619
11.2624 | -34.46 | -18.9711
11.2685 | -34.4605 | -18.9803
11.2746 | -34.4609 | -18.9894
11.2807 | -34.4613 | -18.9985
11.2868 | -34.4618 | -19.0076
11.2929 | -34.4622 | -19.0165
11.299 | -34.4627 | -19.0255
11.3051 | -34.4631 | -19.0346
11.3112 | -34.4635 | -19.0437
11.3173 | -34.464 | -19.0528
11.3234 | -34.4644 | -19.0618
11.3295 | -34.4649 | -19.0708
11.3356 | -34.4653 | -19.08
11.3417 | -34.4657 | -19.0891
11.3478 | -34.4661 | -19.0983
11.3539 | -34.4666 | -19.1075
11.36 | -34.467 | -19.1165
11.3661 | -34.4675 | -19.1255
11.3722 | -34.4679 | -19.1346
11.3783 | -34.4684 | -19.1437
11.3844 | -34.4688 | -19.1529
11.3906 | -34.4693 | -19.1619
11.3967 | -34.4697 | -19.1709
11.4028 | -34.4702 | -19.18
11.4089 | -34.4706 | -19.1892
11.415 | -34.4711 | -19.1982
11.4211 | -34.4715 | -19.2073
11.4272 | -34.472 | -19.2164
11.4333 | -34.4724 | -19.2253
11.4394 | -34.4729 | -19.2344
11.4455 | -34.4733 | -19.2435
11.4516 | -34.4737 | -19.2526
11.4577 | -34.4742 | -19.2616
11.4638 | -34.4747 | -19.2705
11.4699 | -34.4751 | -19.2796
11.476 | -34.4756 | -19.2887
11.4821 | -34.476 | -19.2978
11.4882 | -34.4765 | -19.307
11.4943 | -34.4769 | -19.3161
11.5004 | -34.4774 | -19.325
11.5065 | -34.4778 | -19.3339
11.5126 | -34.4783 | -19.3431
11.5187 | -34.4787 | -19.3522
11.5248 | -34.4792 | -19.3613
11.5309 | -34.4796 | -19.3703
11.5371 | -34.4801 | -19.3793
11.5432 | -34.4805 | -19.3884
11.5493 | -34.481 | -19.3975
11.5554 | -34.4814 | -19.4066
11.5615 | -34.4819 | -19.4156
11.5676 | -34.4823 | -19.4247
11.5737 | -34.4828 | -19.4337
11.5798 | -34.4832 | -19.4427
11.5859 | -34.4837 | -19.4518
11.592 | -34.4841 | -19.4608
11.5981 | -34.4846 | -19.4698
11.6042 | -34.485 | -19.4787
11.6103 | -34.4855 | -19.4878
11.6164 | -34.4859 | -19.4968
11.6225 | -34.4864 | -19.5058
11.6286 | -34.4868 | -19.5147
11.6347 | -34.4872 | -19.5238
11.6408 | -34.4877 | -19.5327
11.6469 | -34.4882 | -19.5417
11.653 | -34.4886 | -19.5508
11.6591 | -34.4891 | -19.5599
11.6652 | -34.4895 | -19.569
11.6713 | -34.49 | -19.578
11.6775 | -34.4904 | -19.5869
11.6836 | -34.4909 | -19.596
11.6897 | -34.4914 | -19.6049
11.6958 | -34.4918 | -19.614
11.7019 | -34.4923 | -19.623
11.708 | -34.4927 | -19.632
11.7141 | -34.4932 | -19.6408
11.7202 | -34.4936 | -19.6498
11.7263 | -34.494 | -19.6589
11.7324 | -34.4945 | -19.6678
11.7385 | -34.495 | -19.6768
11.7446 | -34.4954 | -19.6859
11.7507 | -34.4959 | -19.6947
11.7568 | -34.4963 | -19.7037
11.7629 | -34.4968 | -19.7128
11.769 | -34.4972 | -19.7217
11.7751 | -34.4977 | -19.7307
11.7812 | -34.4981 | -19.7395
11.7873 | -34.4986 | -19.7486
11.7934 | -34.499 | -19.7575
11.7995 | -34.4995 | -19.7665
11.8056 | -34.4999 | -19.7755
11.8117 | -34.5004 | -19.7846
11.8178 | -34.5009 | -19.7934
11.824 | -34.5013 | -19.8024
11.8301 | -34.5018 | -19.8113
11.8362 | -34.5022 | -19.8203
11.8423 | -34.5027 | -19.8293
11.8484 | -34.5031 | -19.8383
11.8545 | -34.5036 | -19.8472
11.8606 | -34.504 | -19.8562
11.8667 | -34.5045 | -19.8652
11.8728 | -34.5049 | -19.874
11.8789 | -34.5054 | -19.8831
11.885 | -34.5058 | -19.892
11.8911 | -34.5063 | -19.9009
11.8972 | -34.5068 | -19.9098
11.9033 | -34.5072 | -19.9188
11.9094 | -34.5077 | -19.9277
11.9155 | -34.5082 | -19.9367
11.9216 | -34.5086 | -19.9454
11.9277 | -34.5091 | -19.9545
11.9338 | -34.5095 | -19.9634
11.9399 | -34.51 | -19.9724
11.946 | -34.5104 | -19.9814
11.9521 | -34.5109 | -19.9902
11.9582 | -34.5114 | -19.999
11.9644 | -34.5118 | -20.0079
11.9705 | -34.5123 | -20.0168
11.9766 | -34.5127 | -20.0257
11.9827 | -34.5132 | -20.0346
11.9888 | -34.5136 | -20.0434
11.9949 | -34.5141 | -20.0523
12.001 | -34.5146 | -20.0612
12.0071 | -34.515 | -20.0702
12.0132 | -34.5155 | -20.0791
12.0193 | -34.5159 | -20.088
12.0254 | -34.5164 | -20.097
12.0315 | -34.5168 | -20.1057
12.0376 | -34.5173 | -20.1145
12.0437 | -34.5177 | -20.1235
12.0498 | -34.5182 | -20.1323
12.0559 | -34.5186 | -20.1413
12.062 | -34.5191 | -20.15
12.0681 | -34.5196 | -20.1589
12.0742 | -34.52 | -20.1677
12.0803 | -34.5205 | -20.1767
12.0864 | -34.5209 | -20.1855
12.0925 | -34.5214 | -20.1944
12.0986 | -34.5219 | -20.2032
12.1047 | -34.5223 | -20.2121
12.1109 | -34.5228 | -20.2209
12.117 | -34.5232 | -20.2299
12.1231 | -34.5237 | -20.2387
12.1292 | -34.5241 | -20.2477
12.1353 | -34.5246 | -20.2565
12.1414 | -34.5251 | -20.2653
12.1475 | -34.5255 | -20.2742
12.1536 | -34.526 | -20.2831
12.1597 | -34.5265 | -20.2919
12.1658 | -34.5269 | -20.3007
12.1719 | -34.5274 | -20.3096
12.178 | -34.5279 | -20.3185
12.1841 | -34.5283 | -20.3274
12.1902 | -34.5288 | -20.3363
12.1963 | -34.5292 | -20.3452
12.2024 | -34.5297 | -20.3539
12.2085 | -34.5302 | -20.3628
12.2146 | -34.5307 | -20.3716
12.2207 | -34.5311 | -20.3805
12.2268 | -34.5316 | -20.3893
12.2329 | -34.532 | -20.398
12.239 | -34.5325 | -20.4067
12.2451 | -34.533 | -20.4155
12.2513 | -34.5334 | -20.4243
12.2574 | -34.5339 | -20.4332
12.2635 | -34.5344 | -20.442
12.2696 | -34.5348 | -20.4509
12.2757 | -34.5353 | -20.4596
12.2818 | -34.5358 | -20.4686
12.2879 | -34.5362 | -20.4774
12.294 | -34.5367 | -20.4863
12.3001 | -34.5371 | -20.4952
12.3062 | -34.5376 | -20.504
12.3123 | -34.5381 | -20.5126
12.3184 | -34.5385 | -20.5215
12.3245 | -34.539 | -20.5303
12.3306 | -34.5395 | -20.5392
12.3367 | -34.5399 | -20.5479
12.3428 | -34.5404 | -20.5567
12.3489 | -34.5408 | -20.5653
12.355 | -34.5413 | -20.5743
12.3611 | -34.5418 | -20.5831
12.3672 | -34.5422 | -20.592
12.3733 | -34.5427 | -20.6007
12.3794 | -34.5432 | -20.6093
12.3855 | -34.5436 | -20.6182
12.3916 | -34.5441 | -20.627
12.3978 | -34.5446 | -20.6358
12.4039 | -34.5451 | -20.6446
12.41 | -34.5455 | -20.6533
12.4161 | -34.546 | -20.6619
12.4222 | -34.5464 | -20.6706
12.4283 | -34.5469 | -20.6795
12.4344 | -34.5474 | -20.6883
12.4405 | -34.5478 | -20.6971
12.4466 | -34.5483 | -20.706
12.4527 | -34.5487 | -20.7147
12.4588 | -34.5492 | -20.7234
12.4649 | -34.5497 | -20.7322
12.471 | -34.5501 | -20.7411
12.4771 | -34.5506 | -20.7499
12.4832 | -34.5511 | -20.7587
12.4893 | -34.5516 | -20.7673
12.4954 | -34.552 | -20.7761
12.5015 | -34.5525 | -20.7849
12.5076 | -34.553 | -20.7937
12.5137 | -34.5534 | -20.8025
12.5198 | -34.5539 | -20.811
12.5259 | -34.5543 | -20.8197
12.532 | -34.5548 | -20.8285
12.5382 | -34.5553 | -20.8372
12.5443 | -34.5557 | -20.8461
12.5504 | -34.5562 | -20.855
12.5565 | -34.5567 | -20.8636
12.5626 | -34.5572 | -20.8724
12.5687 | -34.5576 | -20.8811
12.5748 | -34.5581 | -20.8899
12.5809 | -34.5585 | -20.8986
12.587 | -34.559 | -20.9073
12.5931 | -34.5595 | -20.9158
12.5992 | -34.56 | -20.9245
12.6053 | -34.5605 | -20.9333
12.6114 | -34.561 | -20.9421
12.6175 | -34.5614 | -20.9508
12.6236 | -34.5619 | -20.9596
12.6297 | -34.5624 | -20.9682
12.6358 | -34.5628 | -20.9768
12.6419 | -34.5633 | -20.9856
12.648 | -34.5638 | -20.9943
12.6541 | -34.5642 | -21.003
12.6602 | -34.5647 | -21.0116
12.6663 | -34.5651 | -21.0203
12.6724 | -34.5656 | -21.029
12.6785 | -34.5661 | -21.0377
12.6847 | -34.5665 | -21.0464
12.6908 | -34.567 | -21.0551
12.6969 | -34.5675 | -21.0636
12.703 | -34.5679 | -21.0724
12.7091 | -34.5684 | -21.0811
12.7152 | -34.5688 | -21.0898
12.7213 | -34.5693 | -21.0987
12.7274 | -34.5698 | -21.1073
12.7335 | -34.5703 | -21.1158
12.7396 | -34.5707 | -21.1244
12.7457 | -34.5712 | -21.1332
12.7518 | -34.5717 | -21.1419
12.7579 | -34.5721 | -21.1505
12.764 | -34.5726 | -21.1593
12.7701 | -34.5731 | -21.1679
12.7762 | -34.5735 | -21.1766
12.7823 | -34.574 | -21.1852
12.7884 | -34.5744 | -21.1938
12.7945 | -34.5749 | -21.2027
12.8006 | -34.5754 | -21.2111
12.8067 | -34.5758 | -21.2197
12.8128 | -34.5763 | -21.2283
12.8189 | -34.5767 | -21.237
12.8251 | -34.5772 | -21.2458
12.8312 | -34.5777 | -21.2544
12.8373 | -34.5781 | -21.2629
12.8434 | -34.5786 | -21.2715
12.8495 | -34.5791 | -21.2802
12.8556 | -34.5795 | -21.2889
12.8617 | -34.58 | -21.2976
12.8678 | -34.5805 | -21.3064
12.8739 | -34.581 | -21.3151
12.88 | -34.5814 | -21.3238
12.8861 | -34.5819 | -21.3324
12.8922 | -34.5823 | -21.341
12.8983 | -34.5828 | -21.3498
12.9044 | -34.5833 | -21.3583
12.9105 | -34.5837 | -21.3668
12.9166 | -34.5842 | -21.3756
12.9227 | -34.5846 | -21.3842
12.9288 | -34.5851 | -21.3928
12.9349 | -34.5856 | -21.4014
12.941 | -34.586 | -21.41
12.9471 | -34.5865 | -21.4185
12.9532 | -34.587 | -21.4271
12.9593 | -34.5874 | -21.4358
12.9654 | -34.5879 | -21.4442
12.9716 | -34.5884 | -21.4527
12.9777 | -34.5889 | -21.4614
12.9838 | -34.5893 | -21.47
12.9899 | -34.5898 | -21.4789
12.996 | -34.5902 | -21.4875
13.0021 | -34.5907 | -21.4962
13.0082 | -34.5911 | -21.5047
13.0143 | -34.5916 | -21.513
13.0204 | -34.5921 | -21.5215
13.0265 | -34.5925 | -21.53
13.0326 | -34.593 | -21.5385
13.0387 | -34.5934 | -21.5471
13.0448 | -34.5939 | -21.5558
13.0509 | -34.5944 | -21.5641
13.057 | -34.5948 | -21.5728
13.0631 | -34.5953 | -21.5815
13.0692 | -34.5957 | -21.59
13.0753 | -34.5962 | -21.5987
13.0814 | -34.5967 | -21.6073
13.0875 | -34.5972 | -21.6156
13.0936 | -34.5976 | -21.6242
13.0997 | -34.5981 | -21.6328
13.1058 | -34.5986 | -21.6414
13.112 | -34.599 | -21.6499
13.1181 | -34.5995 | -21.6582
13.1242 | -34.5999 | -21.6668
13.1303 | -34.6004 | -21.6753
13.1364 | -34.6009 | -21.684
13.1425 | -34.6014 | -21.6925
13.1486 | -34.6018 | -21.7009
13.1547 | -34.6023 | -21.7093
13.1608 | -34.6028 | -21.718
13.1669 | -34.6032 | -21.7265
13.173 | -34.6037 | -21.735
13.1791 | -34.6042 | -21.7434
13.1852 | -34.6046 | -21.752
13.1913 | -34.6051 | -21.7605
13.1974 | -34.6055 | -21.7691
13.2035 | -34.606 | -21.7776
13.2096 | -34.6064 | -21.7862
13.2157 | -34.6069 | -21.7946
13.2218 | -34.6074 | -21.803
13.2279 | -34.6078 | -21.8112
13.234 | -34.6083 | -21.8197
13.2401 | -34.6087 | -21.8283
13.2462 | -34.6092 | -21.8369
13.2524 | -34.6097 | -21.8454
13.2585 | -34.6101 | -21.8537
13.2646 | -34.6106 | -21.8622
13.2707 | -34.6111 | -21.8707
13.2768 | -34.6115 | -21.8795
13.2829 | -34.612 | -21.888
13.289 | -34.6125 | -21.8963
13.2951 | -34.613 | -21.9047
13.3012 | -34.6134 | -21.9132
13.3073 | -34.6139 | -21.9216
13.3134 | -34.6144 | -21.9303
13.3195 | -34.6149 | -21.9387
13.3256 | -34.6153 | -21.947
13.3317 | -34.6158 | -21.9548
13.3378 | -34.6163 | -21.9636
13.3439 | -34.6168 | -21.9723
13.35 | -34.6173 | -21.9807
13.3561 | -34.6177 | -21.9893
13.3622 | -34.6182 | -21.9977
13.3683 | -34.6186 | -22.0061
13.3744 | -34.6191 | -22.0146
13.3805 | -34.6195 | -22.023
13.3866 | -34.6199 | -22.0314
13.3927 | -34.6204 | -22.0397
13.3989 | -34.6209 | -22.0484
13.405 | -34.6214 | -22.0572
13.4111 | -34.6218 | -22.0656
13.4172 | -34.6223 | -22.0741
13.4233 | -34.6228 | -22.0828
13.4294 | -34.6232 | -22.0912
13.4355 | -34.6237 | -22.0996
13.4416 | -34.6242 | -22.1078
13.4477 | -34.6247 | -22.1162
13.4538 | -34.6251 | -22.1246
13.4599 | -34.6256 | -22.133
13.466 | -34.6261 | -22.1414
13.4721 | -34.6265 | -22.1493
13.4782 | -34.627 | -22.1578
13.4843 | -34.6275 | -22.1662
13.4904 | -34.628 | -22.1745
13.4965 | -34.6284 | -22.1831
13.5026 | -34.6289 | -22.1919
13.5087 | -34.6294 | -22.2002
13.5148 | -34.6298 | -22.2087
13.5209 | -34.6303 | -22.2173
13.527 | -34.6307 | -22.2259
13.5331 | -34.6312 | -22.2344
13.5393 | -34.6317 | -22.2427
13.5454 | -34.6321 | -22.2509
13.5515 | -34.6326 | -22.2593
13.5576 | -34.633 | -22.2678
13.5637 | -34.6335 | -22.2765
13.5698 | -34.634 | -22.2849
13.5759 | -34.6344 | -22.2928
13.582 | -34.6349 | -22.3012
13.5881 | -34.6353 | -22.3095
13.5942 | -34.6358 | -22.3179
13.6003 | -34.6363 | -22.3263
13.6064 | -34.6368 | -22.3348
13.6125 | -34.6373 | -22.3432
13.6186 | -34.6377 | -22.3516
13.6247 | -34.6382 | -22.3599
13.6308 | -34.6387 | -22.3685
13.6369 | -34.6391 | -22.377
13.643 | -34.6395 | -22.3854
13.6491 | -34.64 | -22.3935
13.6552 | -34.6405 | -22.4018
13.6613 | -34.641 | -22.4103
13.6674 | -34.6415 | -22.4186
13.6735 | -34.6419 | -22.4271
13.6796 | -34.6424 | -22.4354
13.6858 | -34.6429 | -22.4435
13.6919 | -34.6434 | -22.4519
13.698 | -34.6438 | -22.4604
13.7041 | -34.6444 | -22.4687
13.7102 | -34.6448 | -22.4772
13.7163 | -34.6453 | -22.4853
13.7224 | -34.6457 | -22.4936
13.7285 | -34.6462 | -22.502
13.7346 | -34.6467 | -22.5105
13.7407 | -34.6472 | -22.5188
13.7468 | -34.6476 | -22.5272
13.7529 | -34.6481 | -22.5353
13.759 | -34.6486 | -22.5436
13.7651 | -34.649 | -22.5518
13.7712 | -34.6495 | -22.5603
13.7773 | -34.65 | -22.5688
13.7834 | -34.6504 | -22.577
13.7895 | -34.6509 | -22.5851
13.7956 | -34.6514 | -22.5937
13.8017 | -34.6519 | -22.602
13.8078 | -34.6523 | -22.6104
13.8139 | -34.6528 | -22.6187
13.82 | -34.6533 | -22.6271
13.8262 | -34.6538 | -22.6351
13.8323 | -34.6542 | -22.6434
13.8384 | -34.6547 | -22.6516
13.8445 | -34.6552 | -22.66
13.8506 | -34.6556 | -22.6684
13.8567 | -34.6561 | -22.6765
13.8628 | -34.6565 | -22.6848
13.8689 | -34.657 | -22.6931
13.875 | -34.6575 | -22.7015
13.8811 | -34.658 | -22.7099
13.8872 | -34.6584 | -22.7182
13.8933 | -34.6589 | -22.7263
13.8994 | -34.6594 | -22.7347
13.9055 | -34.6598 | -22.7429
13.9116 | -34.6603 | -22.7513
13.9177 | -34.6607 | -22.7596
13.9238 | -34.6612 | -22.768
13.9299 | -34.6617 | -22.776
13.936 | -34.6621 | -22.7845
13.9421 | -34.6626 | -22.7927
13.9482 | -34.663 | -22.801
13.9543 | -34.6635 | -22.8094
13.9604 | -34.6639 | -22.8178
13.9665 | -34.6645 | -22.8258
13.9727 | -34.6649 | -22.8342
13.9788 | -34.6654 | -22.8426
13.9849 | -34.6659 | -22.8509
13.991 | -34.6664 | -22.8592
13.9971 | -34.6668 | -22.8673
14.0032 | -34.6673 | -22.8757
14.0093 | -34.6677 | -22.8839
14.0154 | -34.6682 | -22.8922
14.0215 | -34.6687 | -22.9004
14.0276 | -34.6691 | -22.9085
14.0337 | -34.6696 | -22.9166
14.0398 | -34.6701 | -22.9248
14.0459 | -34.6706 | -22.933
14.052 | -34.671 | -22.9413
14.0581 | -34.6715 | -22.9495
14.0642 | -34.672 | -22.9578
14.0703 | -34.6725 | -22.9659
14.0764 | -34.6729 | -22.9742
14.0825 | -34.6734 | -22.9825
14.0886 | -34.6738 | -22.9905
14.0947 | -34.6743 | -22.9987
14.1008 | -34.6748 | -23.0071
14.1069 | -34.6752 | -23.0153
14.1131 | -34.6757 | -23.0238
14.1192 | -34.6762 | -23.0319
14.1253 | -34.6766 | -23.0401
14.1314 | -34.6771 | -23.0483
14.1375 | -34.6776 | -23.0563
14.1436 | -34.678 | -23.0645
14.1497 | -34.6785 | -23.0727
14.1558 | -34.6789 | -23.081
14.1619 | -34.6794 | -23.0894
14.168 | -34.6799 | -23.0976
14.1741 | -34.6803 | -23.1056
14.1802 | -34.6808 | -23.1138
14.1863 | -34.6813 | -23.1221
14.1924 | -34.6818 | -23.1303
14.1985 | -34.6822 | -23.1384
14.2046 | -34.6827 | -23.1466
14.2107 | -34.6832 | -23.1546
14.2168 | -34.6837 | -23.1629
14.2229 | -34.6842 | -23.1711
14.229 | -34.6847 | -23.1793
14.2351 | -34.6851 | -23.1875
14.2412 | -34.6856 | -23.1959
14.2473 | -34.6861 | -23.2037
14.2534 | -34.6866 | -23.2118
14.2596 | -34.687 | -23.22
14.2657 | -34.6875 | -23.2282
14.2718 | -34.688 | -23.2364
14.2779 | -34.6885 | -23.2447
14.284 | -34.6889 | -23.2527
14.2901 | -34.6894 | -23.2609
14.2962 | -34.6899 | -23.2692
14.3023 | -34.6904 | -23.2773
14.3084 | -34.6909 | -23.2856
14.3145 | -34.6914 | -23.2936
14.3206 | -34.6918 | -23.3019
14.3267 | -34.6923 | -23.3101
14.3328 | -34.6928 | -23.3184
14.3389 | -34.6933 | -23.3267
14.345 | -34.6937 | -23.3349
14.3511 | -34.6942 | -23.3429
14.3572 | -34.6947 | -23.3511
14.3633 | -34.6952 | -23.3593
14.3694 | -34.6956 | -23.3675
14.3755 | -34.6961 | -23.3757
14.3816 | -34.6966 | -23.3839
14.3877 | -34.6971 | -23.392
14.3938 | -34.6976 | -23.4002
14.4 | -34.6981 | -23.4085
14.4061 | -34.6985 | -23.4166
14.4122 | -34.699 | -23.4249
14.4183 | -34.6995 | -23.433
14.4244 | -34.7 | -23.4411
14.4305 | -34.7005 | -23.4493
14.4366 | -34.701 | -23.4575
14.4427 | -34.7014 | -23.4657
14.4488 | -34.7019 | -23.4738
14.4549 | -34.7024 | -23.4818
14.461 | -34.7029 | -23.49
14.4671 | -34.7034 | -23.4982
14.4732 | -34.7039 | -23.5065
14.4793 | -34.7044 | -23.5145
14.4854 | -34.7049 | -23.5226
14.4915 | -34.7053 | -23.5307
14.4976 | -34.7058 | -23.5388
14.5037 | -34.7063 | -23.5469
14.5098 | -34.7068 | -23.5551
14.5159 | -34.7073 | -23.5633
14.522 | -34.7077 | -23.5715
14.5281 | -34.7082 | -23.5796
14.5342 | -34.7087 | -23.5878
14.5403 | -34.7092 | -23.5959
14.5465 | -34.7097 | -23.6041
14.5526 | -34.7102 | -23.6123
14.5587 | -34.7107 | -23.6204
14.5648 | -34.7112 | -23.6282
14.5709 | -34.7116 | -23.6362
14.577 | -34.7121 | -23.6444
14.5831 | -34.7126 | -23.6525
14.5892 | -34.7131 | -23.6606
14.5953 | -34.7136 | -23.6686
14.6014 | -34.714 | -23.6768
14.6075 | -34.7145 | -23.685
14.6136 | -34.715 | -23.6934
14.6197 | -34.7155 | -23.7016
14.6258 | -34.716 | -23.7097
14.6319 | -34.7165 | -23.7175
14.638 | -34.7169 | -23.7256
14.6441 | -34.7174 | -23.7339
14.6502 | -34.7179 | -23.7421
14.6563 | -34.7184 | -23.7503
14.6624 | -34.7189 | -23.7584
14.6685 | -34.7193 | -23.7662
14.6746 | -34.7198 | -23.7744
14.6807 | -34.7203 | -23.7825
14.6869 | -34.7207 | -23.7907
14.693 | -34.7212 | -23.7987
14.6991 | -34.7217 | -23.8069
14.7052 | -34.7222 | -23.8147
14.7113 | -34.7227 | -23.8227
14.7174 | -34.7232 | -23.8309
14.7235 | -34.7237 | -23.839
14.7296 | -34.7242 | -23.8472
14.7357 | -34.7247 | -23.8552
14.7418 | -34.7252 | -23.8633
14.7479 | -34.7256 | -23.8714
14.754 | -34.7261 | -23.8795
14.7601 | -34.7266 | -23.8876
14.7662 | -34.7271 | -23.8957
14.7723 | -34.7275 | -23.9037
14.7784 | -34.728 | -23.9119
14.7845 | -34.7285 | -23.92
14.7906 | -34.729 | -23.928
14.7967 | -34.7295 | -23.9361
14.8028 | -34.73 | -23.9441
14.8089 | -34.7305 | -23.9522
14.815 | -34.7309 | -23.9603
14.8211 | -34.7314 | -23.9684
14.8272 | -34.7319 | -23.9764
14.8334 | -34.7324 | -23.9844
14.8395 | -34.7329 | -23.9926
14.8456 | -34.7334 | -24.0006
14.8517 | -34.7339 | -24.0088
14.8578 | -34.7344 | -24.0169
14.8639 | -34.7349 | -24.0249
14.87 | -34.7354 | -24.0328
14.8761 | -34.7359 | -24.0407
14.8822 | -34.7364 | -24.0488
14.8883 | -34.7369 | -24.0568
14.8944 | -34.7373 | -24.0647
14.9005 | -34.7378 | -24.0729
14.9066 | -34.7383 | -24.0809
14.9127 | -34.7388 | -24.0888
14.9188 | -34.7392 | -24.0968
14.9249 | -34.7397 | -24.1049
14.931 | -34.7402 | -24.1129
14.9371 | -34.7407 | -24.121
14.9432 | -34.7412 | -24.1291
14.9493 | -34.7417 | -24.137
14.9554 | -34.7422 | -24.145
14.9615 | -34.7427 | -24.1531
14.9676 | -34.7432 | -24.1613
14.9738 | -34.7437 | -24.1692
14.9799 | -34.7442 | -24.1773
14.986 | -34.7447 | -24.1852
14.9921 | -34.7452 | -24.1933
14.9982 | -34.7457 | -24.2013
15.0043 | -34.7461 | -24.2092
15.0104 | -34.7466 | -24.2172
15.0165 | -34.7471 | -24.2252
15.0226 | -34.7477 | -24.2329
15.0287 | -34.7482 | -24.241
15.0348 | -34.7487 | -24.249
15.0409 | -34.7492 | -24.257
15.047 | -34.7497 | -24.2651
15.0531 | -34.7502 | -24.2729
15.0592 | -34.7506 | -24.281
15.0653 | -34.7511 | -24.289
15.0714 | -34.7516 | -24.297
15.0775 | -34.7521 | -24.305
15.0836 | -34.7526 | -24.3129
15.0897 | -34.7531 | -24.3209
15.0958 | -34.7536 | -24.3289
15.1019 | -34.7541 | -24.337
15.108 | -34.7546 | -24.345
15.1141 | -34.7551 | -24.3532
15.1203 | -34.7556 | -24.3613
15.1264 | -34.7561 | -24.3691
15.1325 | -34.7566 | -24.3772
15.1386 | -34.7571 | -24.3851
15.1447 | -34.7576 | -24.3931
15.1508 | -34.7581 | -24.401
15.1569 | -34.7586 | -24.409
15.163 | -34.7591 | -24.4169
15.1691 | -34.7596 | -24.4248
15.1752 | -34.7601 | -24.4328
15.1813 | -34.7606 | -24.4408
15.1874 | -34.7611 | -24.4487
15.1935 | -34.7616 | -24.4565
15.1996 | -34.7621 | -24.4644
15.2057 | -34.7626 | -24.4725
15.2118 | -34.7631 | -24.4806
15.2179 | -34.7636 | -24.4885
15.224 | -34.7641 | -24.4966
15.2301 | -34.7646 | -24.5043
15.2362 | -34.7651 | -24.5123
15.2423 | -34.7656 | -24.5203
15.2484 | -34.7661 | -24.5282
15.2545 | -34.7666 | -24.5361
15.2607 | -34.7671 | -24.5441
15.2668 | -34.7676 | -24.552
15.2729 | -34.7681 | -24.56
15.279 | -34.7686 | -24.568
15.2851 | -34.7691 | -24.5761
15.2912 | -34.7696 | -24.584
15.2973 | -34.7701 | -24.592
15.3034 | -34.7706 | -24.5998
15.3095 | -34.7711 | -24.6076
15.3156 | -34.7715 | -24.6155
15.3217 | -34.772 | -24.6236
15.3278 | -34.7726 | -24.6316
15.3339 | -34.7731 | -24.6394
15.34 | -34.7736 | -24.6473
15.3461 | -34.7741 | -24.6554
15.3522 | -34.7746 | -24.6633
15.3583 | -34.7751 | -24.6713
15.3644 | -34.7756 | -24.6793
15.3705 | -34.7761 | -24.6869
15.3766 | -34.7766 | -24.6947
15.3827 | -34.7771 | -24.7027
15.3888 | -34.7776 | -24.7106
15.3949 | -34.7781 | -24.7186
15.401 | -34.7786 | -24.7266
15.4072 | -34.7791 | -24.7344
15.4133 | -34.7796 | -24.7424
15.4194 | -34.7801 | -24.7504
15.4255 | -34.7806 | -24.7584
15.4316 | -34.7811 | -24.7662
15.4377 | -34.7816 | -24.7742
15.4438 | -34.7821 | -24.7818
15.4499 | -34.7826 | -24.7898
15.456 | -34.7831 | -24.7977
15.4621 | -34.7836 | -24.8056
15.4682 | -34.7841 | -24.8135
15.4743 | -34.7846 | -24.8212
15.4804 | -34.7851 | -24.8292
15.4865 | -34.7856 | -24.8371
15.4926 | -34.7861 | -24.8451
15.4987 | -34.7866 | -24.8529
15.5048 | -34.7871 | -24.8608
15.5109 | -34.7877 | -24.8686
15.517 | -34.7882 | -24.8766
15.5231 | -34.7887 | -24.8846
15.5292 | -34.7892 | -24.8926
15.5353 | -34.7897 | -24.9007
15.5414 | -34.7901 | -24.9085
15.5476 | -34.7907 | -24.9162
15.5537 | -34.7912 | -24.9241
15.5598 | -34.7916 | -24.932
15.5659 | -34.7921 | -24.9399
15.572 | -34.7926 | -24.9478
15.5781 | -34.7931 | -24.9558
15.5842 | -34.7936 | -24.9638
15.5903 | -34.7941 | -24.9717
15.5964 | -34.7946 | -24.9796
15.6025 | -34.7951 | -24.9873
15.6086 | -34.7957 | -24.9951
15.6147 | -34.7962 | -25.0031
15.6208 | -34.7967 | -25.0108
15.6269 | -34.7972 | -25.0186
15.633 | -34.7977 | -25.0264
15.6391 | -34.7982 | -25.0343
15.6452 | -34.7987 | -25.0421
15.6513 | -34.7992 | -25.0499
15.6574 | -34.7997 | -25.0579
15.6635 | -34.8002 | -25.0658
15.6696 | -34.8007 | -25.0736
15.6757 | -34.8012 | -25.0815
15.6818 | -34.8017 | -25.0894
15.688 | -34.8022 | -25.0972
15.6941 | -34.8027 | -25.1052
15.7002 | -34.8032 | -25.1132
15.7063 | -34.8037 | -25.121
15.7124 | -34.8042 | -25.129
15.7185 | -34.8047 | -25.1368
15.7246 | -34.8053 | -25.1445
15.7307 | -34.8058 | -25.1522
15.7368 | -34.8063 | -25.16
15.7429 | -34.8068 | -25.1677
15.749 | -34.8073 | -25.1756
15.7551 | -34.8078 | -25.1835
15.7612 | -34.8083 | -25.1913
15.7673 | -34.8088 | -25.1991
15.7734 | -34.8093 | -25.207
15.7795 | -34.8098 | -25.2149
15.7856 | -34.8103 | -25.2227
15.7917 | -34.8108 | -25.2303
15.7978 | -34.8113 | -25.2382
15.8039 | -34.8118 | -25.2461
15.81 | -34.8124 | -25.254
15.8161 | -34.8129 | -25.262
15.8222 | -34.8134 | -25.2699
15.8283 | -34.8139 | -25.2777
15.8345 | -34.8144 | -25.2855
15.8406 | -34.8149 | -25.2933
15.8467 | -34.8154 | -25.3011
15.8528 | -34.8159 | -25.3089
15.8589 | -34.8164 | -25.3166
15.865 | -34.8169 | -25.3242
15.8711 | -34.8174 | -25.3321
15.8772 | -34.8179 | -25.3398
15.8833 | -34.8184 | -25.3476
15.8894 | -34.8189 | -25.3555
15.8955 | -34.8194 | -25.3634
15.9016 | -34.8199 | -25.371
15.9077 | -34.8204 | -25.3788
15.9138 | -34.8209 | -25.3867
15.9199 | -34.8214 | -25.3944
15.926 | -34.8219 | -25.4023
15.9321 | -34.8224 | -25.4099
15.9382 | -34.823 | -25.4177
15.9443 | -34.8235 | -25.4254
15.9504 | -34.824 | -25.4333
15.9565 | -34.8245 | -25.4412
15.9626 | -34.825 | -25.4491
15.9687 | -34.8255 | -25.4567
15.9749 | -34.826 | -25.4643
15.981 | -34.8265 | -25.4721
15.9871 | -34.8271 | -25.48
15.9932 | -34.8276 | -25.4878
15.9993 | -34.8281 | -25.4956
16.0054 | -34.8286 | -25.5034
16.0115 | -34.8291 | -25.5111
16.0176 | -34.8297 | -25.519
16.0237 | -34.8302 | -25.5269
16.0298 | -34.8307 | -25.5347
16.0359 | -34.8312 | -25.5427
16.042 | -34.8317 | -25.5505
16.0481 | -34.8322 | -25.5583
16.0542 | -34.8327 | -25.5662
16.0603 | -34.8332 | -25.574
16.0664 | -34.8337 | -25.5819
16.0725 | -34.8342 | -25.5893
16.0786 | -34.8348 | -25.597
16.0847 | -34.8353 | -25.6048
16.0908 | -34.8358 | -25.6126
16.0969 | -34.8363 | -25.6204
16.103 | -34.8368 | -25.6282
16.1091 | -34.8374 | -25.6358
16.1152 | -34.8379 | -25.6436
16.1214 | -34.8384 | -25.6515
16.1275 | -34.8389 | -25.6593
16.1336 | -34.8394 | -25.6672
16.1397 | -34.84 | -25.675
16.1458 | -34.8404 | -25.6827
16.1519 | -34.841 | -25.6904
16.158 | -34.8415 | -25.6983
16.1641 | -34.842 | -25.706
16.1702 | -34.8425 | -25.7137
16.1763 | -34.843 | -25.7215
16.1824 | -34.8435 | -25.7292
16.1885 | -34.844 | -25.7372
16.1946 | -34.8446 | -25.7448
16.2007 | -34.8451 | -25.7525
16.2068 | -34.8456 | -25.7601
16.2129 | -34.8461 | -25.7678
16.219 | -34.8467 | -25.7755
16.2251 | -34.8472 | -25.7833
16.2312 | -34.8477 | -25.791
16.2373 | -34.8482 | -25.7988
16.2434 | -34.8487 | -25.8066
16.2495 | -34.8492 | -25.8142
16.2556 | -34.8497 | -25.8221
16.2618 | -34.8502 | -25.83
16.2679 | -34.8507 | -25.8378
16.274 | -34.8512 | -25.8455
16.2801 | -34.8517 | -25.8534
16.2862 | -34.8522 | -25.8609
16.2923 | -34.8527 | -25.8685
16.2984 | -34.8532 | -25.8762
16.3045 | -34.8537 | -25.884
16.3106 | -34.8542 | -25.8918
16.3167 | -34.8547 | -25.8993
16.3228 | -34.8553 | -25.9067
16.3289 | -34.8558 | -25.9143
16.335 | -34.8563 | -25.9221
16.3411 | -34.8568 | -25.9299
16.3472 | -34.8572 | -25.9378
16.3533 | -34.8578 | -25.9454
16.3594 | -34.8583 | -25.953
16.3655 | -34.8588 | -25.9606
16.3716 | -34.8593 | -25.9684
16.3777 | -34.8598 | -25.9762
16.3838 | -34.8603 | -25.984
16.3899 | -34.8608 | -25.9915
16.396 | -34.8613 | -25.9993
16.4021 | -34.8619 | -26.007
16.4083 | -34.8623 | -26.0146
16.4144 | -34.8629 | -26.0223
16.4205 | -34.8634 | -26.0302
16.4266 | -34.8639 | -26.0379
16.4327 | -34.8644 | -26.0456
16.4388 | -34.8648 | -26.0532
16.4449 | -34.8653 | -26.0608
16.451 | -34.8658 | -26.0687
16.4571 | -34.8663 | -26.0765
16.4632 | -34.8668 | -26.0839
16.4693 | -34.8673 | -26.0916
16.4754 | -34.8678 | -26.0993
16.4815 | -34.8683 | -26.1069
16.4876 | -34.8688 | -26.1143
16.4937 | -34.8693 | -26.122
16.4998 | -34.8698 | -26.1295
16.5059 | -34.8703 | -26.137
16.512 | -34.8708 | -26.1449
16.5181 | -34.8714 | -26.1525
16.5242 | -34.8719 | -26.1603
16.5303 | -34.8724 | -26.1677
16.5364 | -34.8729 | -26.1755
16.5425 | -34.8734 | -26.1833
16.5487 | -34.8739 | -26.1908
16.5548 | -34.8744 | -26.1982
16.5609 | -34.8749 | -26.2059
16.567 | -34.8754 | -26.2136
16.5731 | -34.8759 | -26.2212
16.5792 | -34.8764 | -26.2291
16.5853 | -34.8769 | -26.2366
16.5914 | -34.8774 | -26.2444
16.5975 | -34.8779 | -26.2521
16.6036 | -34.8785 | -26.2595
16.6097 | -34.8789 | -26.2674
16.6158 | -34.8795 | -26.275
16.6219 | -34.88 | -26.2825
16.628 | -34.8805 | -26.2902
16.6341 | -34.881 | -26.298
16.6402 | -34.8816 | -26.3057
16.6463 | -34.8821 | -26.3133
16.6524 | -34.8826 | -26.3211
16.6585 | -34.8831 | -26.3289
16.6646 | -34.8837 | -26.3367
16.6707 | -34.8842 | -26.3444
16.6768 | -34.8846 | -26.3521
16.6829 | -34.8851 | -26.3595
16.689 | -34.8857 | -26.3673
16.6952 | -34.8862 | -26.375
16.7013 | -34.8867 | -26.3826
16.7074 | -34.8872 | -26.3901
16.7135 | -34.8878 | -26.3979
16.7196 | -34.8883 | -26.4058
16.7257 | -34.8888 | -26.4134
16.7318 | -34.8893 | -26.4211
16.7379 | -34.8898 | -26.4287
16.744 | -34.8903 | -26.4362
16.7501 | -34.8907 | -26.4437
16.7562 | -34.8912 | -26.4511
16.7623 | -34.8917 | -26.4588
16.7684 | -34.8922 | -26.4665
16.7745 | -34.8928 | -26.4739
16.7806 | -34.8932 | -26.4813
16.7867 | -34.8938 | -26.4889
16.7928 | -34.8943 | -26.4967
16.7989 | -34.8948 | -26.5044
16.805 | -34.8953 | -26.5121
16.8111 | -34.8957 | -26.5194
16.8172 | -34.8962 | -26.5272
16.8233 | -34.8968 | -26.5349
16.8294 | -34.8973 | -26.5426
16.8356 | -34.8978 | -26.5502
16.8417 | -34.8983 | -26.5577
16.8478 | -34.8988 | -26.5654
16.8539 | -34.8993 | -26.573
16.86 | -34.8998 | -26.5806
16.8661 | -34.9003 | -26.5882
16.8722 | -34.9008 | -26.5961
16.8783 | -34.9013 | -26.6038
16.8844 | -34.9018 | -26.6114
16.8905 | -34.9023 | -26.6192
16.8966 | -34.9028 | -26.6267
16.9027 | -34.9033 | -26.6342
16.9088 | -34.9038 | -26.6418
16.9149 | -34.9043 | -26.6494
16.921 | -34.9048 | -26.6568
16.9271 | -34.9053 | -26.6642
16.9332 | -34.9058 | -26.6719
16.9393 | -34.9063 | -26.6797
16.9454 | -34.9068 | -26.6872
16.9515 | -34.9073 | -26.6949
16.9576 | -34.9078 | -26.7021
16.9637 | -34.9082 | -26.7099
16.9698 | -34.9087 | -26.7176
16.9759 | -34.9092 | -26.7252
16.9821 | -34.9097 | -26.7328
16.9882 | -34.9102 | -26.7404
16.9943 | -34.9107 | -26.748
17.0004 | -34.9112 | -26.7555
17.0065 | -34.9117 | -26.763
17.0126 | -34.9122 | -26.7707
17.0187 | -34.9127 | -26.7784
17.0248 | -34.9132 | -26.7862
17.0309 | -34.9138 | -26.7937
17.037 | -34.9143 | -26.8014
17.0431 | -34.9148 | -26.8094
17.0492 | -34.9153 | -26.8172
17.0553 | -34.9158 | -26.8251
17.0614 | -34.9163 | -26.8324
17.0675 | -34.9168 | -26.8397
17.0736 | -34.9173 | -26.8469
17.0797 | -34.9177 | -26.8547
17.0858 | -34.9182 | -26.8625
17.0919 | -34.9187 | -26.8703
17.098 | -34.9191 | -26.8774
17.1041 | -34.9196 | -26.8852
17.1102 | -34.9201 | -26.8929
17.1163 | -34.9206 | -26.9007
17.1225 | -34.9211 | -26.9084
17.1286 | -34.9216 | -26.9157
17.1347 | -34.9221 | -26.9233
17.1408 | -34.9226 | -26.931
17.1469 | -34.9231 | -26.9384
17.153 | -34.9236 | -26.9461
17.1591 | -34.9241 | -26.9537
17.1652 | -34.9246 | -26.9613
17.1713 | -34.9251 | -26.9687
17.1774 | -34.9256 | -26.9763
17.1835 | -34.9261 | -26.984
17.1896 | -34.9265 | -26.9914
17.1957 | -34.927 | -26.999
17.2018 | -34.9275 | -27.0066
17.2079 | -34.928 | -27.0139
17.214 | -34.9285 | -27.0216
17.2201 | -34.9289 | -27.0293
17.2262 | -34.9294 | -27.037
17.2323 | -34.9299 | -27.0445
17.2384 | -34.9303 | -27.0522
17.2445 | -34.9308 | -27.0598
17.2506 | -34.9313 | -27.0675
17.2567 | -34.9318 | -27.0751
17.2628 | -34.9323 | -27.0828
17.269 | -34.9328 | -27.0902
17.2751 | -34.9333 | -27.0977
17.2812 | -34.9338 | -27.1054
17.2873 | -34.9342 | -27.1131
17.2934 | -34.9347 | -27.1207
17.2995 | -34.9352 | -27.1282
17.3056 | -34.9357 | -27.1357
17.3117 | -34.9362 | -27.1434
17.3178 | -34.9367 | -27.151
17.3239 | -34.9372 | -27.1586
17.33 | -34.9377 | -27.1662
17.3361 | -34.9382 | -27.1738
17.3422 | -34.9386 | -27.1807
17.3483 | -34.9391 | -27.1883
17.3544 | -34.9396 | -27.1959
17.3605 | -34.9401 | -27.2036
17.3666 | -34.9406 | -27.2113
17.3727 | -34.9411 | -27.219
17.3788 | -34.9416 | -27.2263
17.3849 | -34.9421 | -27.234
17.391 | -34.9426 | -27.2416
17.3971 | -34.9431 | -27.2491
17.4032 | -34.9436 | -27.2568
17.4094 | -34.944 | -27.2642
17.4155 | -34.9445 | -27.2716
17.4216 | -34.9449 | -27.279
17.4277 | -34.9455 | -27.2866
17.4338 | -34.9459 | -27.2941
17.4399 | -34.9464 | -27.3013
17.446 | -34.9468 | -27.3086
17.4521 | -34.9473 | -27.3159
17.4582 | -34.9478 | -27.3236
17.4643 | -34.9483 | -27.3314
17.4704 | -34.9488 | -27.3389
17.4765 | -34.9493 | -27.3464
17.4826 | -34.9497 | -27.3539
17.4887 | -34.9502 | -27.3616
17.4948 | -34.9507 | -27.3691
17.5009 | -34.9512 | -27.3765
17.507 | -34.9517 | -27.3842
17.5131 | -34.9521 | -27.3921
17.5192 | -34.9527 | -27.3994
17.5253 | -34.9532 | -27.4071
17.5314 | -34.9537 | -27.4146
17.5375 | -34.9542 | -27.4222
17.5436 | -34.9547 | -27.43
17.5497 | -34.9551 | -27.4375
17.5559 | -34.9556 | -27.4447
17.562 | -34.9561 | -27.452
17.5681 | -34.9566 | -27.4592
17.5742 | -34.9571 | -27.467
17.5803 | -34.9575 | -27.4747
17.5864 | -34.958 | -27.482
17.5925 | -34.9584 | -27.4894
17.5986 | -34.9589 | -27.4969
17.6047 | -34.9594 | -27.5044
17.6108 | -34.9598 | -27.5121
17.6169 | -34.9603 | -27.5198
17.623 | -34.9608 | -27.5272
17.6291 | -34.9613 | -27.5349
17.6352 | -34.9617 | -27.5425
17.6413 | -34.9622 | -27.5501
17.6474 | -34.9627 | -27.5574
17.6535 | -34.9632 | -27.5647
17.6596 | -34.9636 | -27.5719
17.6657 | -34.9641 | -27.5793
17.6718 | -34.9645 | -27.5865
17.6779 | -34.965 | -27.5939
17.684 | -34.9655 | -27.6016
17.6901 | -34.966 | -27.6089
17.6963 | -34.9664 | -27.6163
17.7024 | -34.9669 | -27.6239
17.7085 | -34.9674 | -27.6314
17.7146 | -34.9678 | -27.639
17.7207 | -34.9682 | -27.6465
17.7268 | -34.9686 | -27.6537
17.7329 | -34.9691 | -27.6611
17.739 | -34.9696 | -27.6685
17.7451 | -34.9701 | -27.6761
17.7512 | -34.9706 | -27.6837
17.7573 | -34.9711 | -27.691
17.7634 | -34.9715 | -27.6982
17.7695 | -34.972 | -27.7053
17.7756 | -34.9725 | -27.7127
17.7817 | -34.973 | -27.7202
17.7878 | -34.9735 | -27.7276
17.7939 | -34.9739 | -27.7352
17.8 | -34.9744 | -27.7426
17.8061 | -34.9748 | -27.7499
17.8122 | -34.9753 | -27.7577
17.8183 | -34.9757 | -27.7653
17.8244 | -34.9762 | -27.7728
17.8305 | -34.9767 | -27.7802
17.8366 | -34.9771 | -27.7877
17.8428 | -34.9776 | -27.7952
17.8489 | -34.978 | -27.8025
17.855 | -34.9785 | -27.8101
17.8611 | -34.979 | -27.8177
17.8672 | -34.9795 | -27.8252
17.8733 | -34.98 | -27.8329
17.8794 | -34.9804 | -27.8406
17.8855 | -34.9809 | -27.8482
17.8916 | -34.9813 | -27.8558
17.8977 | -34.9818 | -27.8633
17.9038 | -34.9822 | -27.8709
17.9099 | -34.9827 | -27.8786
17.916 | -34.9831 | -27.8858
17.9221 | -34.9836 | -27.8934
17.9282 | -34.9841 | -27.9008
17.9343 | -34.9846 | -27.9084
17.9404 | -34.985 | -27.916
17.9465 | -34.9855 | -27.9235
17.9526 | -34.9859 | -27.9309
17.9587 | -34.9864 | -27.9385
17.9648 | -34.9869 | -27.9458
17.9709 | -34.9873 | -27.9534
17.977 | -34.9878 | -27.9606
17.9832 | -34.9883 | -27.968
17.9893 | -34.9887 | -27.9755
17.9954 | -34.9892 | -27.9827
18.0015 | -34.9896 | -27.9902
18.0076 | -34.9901 | -27.9977
18.0137 | -34.9906 | -28.0052
18.0198 | -34.991 | -28.0127
18.0259 | -34.9915 | -28.0199
18.032 | -34.9919 | -28.0274
18.0381 | -34.9924 | -28.0347
18.0442 | -34.9928 | -28.0419
18.0503 | -34.9932 | -28.0494
18.0564 | -34.9937 | -28.0567
18.0625 | -34.9942 | -28.0641
18.0686 | -34.9946 | -28.0714
18.0747 | -34.9951 | -28.079
18.0808 | -34.9955 | -28.0862
18.0869 | -34.996 | -28.0939
18.093 | -34.9964 | -28.1009
18.0991 | -34.9969 | -28.1084
18.1052 | -34.9973 | -28.116
18.1113 | -34.9978 | -28.1234
18.1174 | -34.9983 | -28.1307
18.1236 | -34.9987 | -28.1381
18.1297 | -34.9991 | -28.1453
18.1358 | -34.9996 | -28.1526
18.1419 | -35 | -28.16
18.148 | -35.0005 | -28.1676
18.1541 | -35.001 | -28.1746
18.1602 | -35.0015 | -28.182
18.1663 | -35.0019 | -28.1894
18.1724 | -35.0024 | -28.1968
18.1785 | -35.0028 | -28.2042
18.1846 | -35.0032 | -28.2118
18.1907 | -35.0037 | -28.219
18.1968 | -35.0041 | -28.2265
18.2029 | -35.0047 | -28.234
18.209 | -35.0052 | -28.2414
18.2151 | -35.0056 | -28.2491
18.2212 | -35.0061 | -28.256
18.2273 | -35.0065 | -28.2634
18.2334 | -35.007 | -28.2706
18.2395 | -35.0075 | -28.2779
18.2456 | -35.0079 | -28.2854
18.2517 | -35.0083 | -28.2926
18.2578 | -35.0088 | -28.3
18.2639 | -35.0092 | -28.3074
18.2701 | -35.0097 | -28.3144
18.2762 | -35.0102 | -28.3219
18.2823 | -35.0106 | -28.3294
18.2884 | -35.0111 | -28.3368
18.2945 | -35.0115 | -28.344
18.3006 | -35.0119 | -28.3514
18.3067 | -35.0123 | -28.3589
18.3128 | -35.0128 | -28.3664
18.3189 | -35.0132 | -28.3739
18.325 | -35.0137 | -28.3812
18.3311 | -35.0141 | -28.3884
18.3372 | -35.0146 | -28.3959
18.3433 | -35.015 | -28.4031
18.3494 | -35.0155 | -28.4103
18.3555 | -35.0159 | -28.4175
18.3616 | -35.0164 | -28.4249
18.3677 | -35.0169 | -28.4322
18.3738 | -35.0173 | -28.4397
18.3799 | -35.0177 | -28.4471
18.386 | -35.0182 | -28.4545
18.3921 | -35.0186 | -28.4619
18.3982 | -35.019 | -28.4693
18.4043 | -35.0195 | -28.4767
18.4105 | -35.02 | -28.4843
18.4166 | -35.0204 | -28.4915
18.4227 | -35.0209 | -28.499
18.4288 | -35.0213 | -28.5064
18.4349 | -35.0218 | -28.5132
18.441 | -35.0222 | -28.5208
18.4471 | -35.0227 | -28.5283
18.4532 | -35.0232 | -28.5354
18.4593 | -35.0236 | -28.5426
18.4654 | -35.0241 | -28.5498
18.4715 | -35.0244 | -28.5573
18.4776 | -35.0249 | -28.5647
18.4837 | -35.0254 | -28.5722
18.4898 | -35.0258 | -28.5796
18.4959 | -35.0263 | -28.5871
18.502 | -35.0267 | -28.5945
18.5081 | -35.0272 | -28.6022
18.5142 | -35.0276 | -28.6096
18.5203 | -35.028 | -28.6171
18.5264 | -35.0285 | -28.6244
18.5325 | -35.0289 | -28.6315
18.5386 | -35.0293 | -28.639
18.5447 | -35.0298 | -28.6463
18.5508 | -35.0302 | -28.6536
18.557 | -35.0307 | -28.6609
18.5631 | -35.0311 | -28.6681
18.5692 | -35.0315 | -28.6754
18.5753 | -35.032 | -28.6828
18.5814 | -35.0324 | -28.6899
18.5875 | -35.0329 | -28.6971
18.5936 | -35.0333 | -28.7044
18.5997 | -35.0338 | -28.7121
18.6058 | -35.0342 | -28.7193
18.6119 | -35.0346 | -28.7264
18.618 | -35.035 | -28.7339
18.6241 | -35.0355 | -28.7412
18.6302 | -35.0359 | -28.7485
18.6363 | -35.0364 | -28.7559
18.6424 | -35.0368 | -28.7634
18.6485 | -35.0373 | -28.7706
18.6546 | -35.0377 | -28.778
18.6607 | -35.0381 | -28.7854
18.6668 | -35.0385 | -28.7926
18.6729 | -35.039 | -28.8001
18.679 | -35.0394 | -28.8072
18.6851 | -35.0398 | -28.8147
18.6912 | -35.0402 | -28.8219
18.6974 | -35.0406 | -28.8291
18.7035 | -35.041 | -28.8364
18.7096 | -35.0415 | -28.8438
18.7157 | -35.042 | -28.851
18.7218 | -35.0424 | -28.8584
18.7279 | -35.0428 | -28.8656
18.734 | -35.0433 | -28.8727
18.7401 | -35.0437 | -28.8805
18.7462 | -35.0442 | -28.8879
18.7523 | -35.0446 | -28.8952
18.7584 | -35.0451 | -28.9025
18.7645 | -35.0455 | -28.9098
18.7706 | -35.0459 | -28.917
18.7767 | -35.0464 | -28.9246
18.7828 | -35.0468 | -28.9318
18.7889 | -35.0472 | -28.9392
18.795 | -35.0476 | -28.9465
18.8011 | -35.0481 | -28.9537
18.8072 | -35.0486 | -28.9612
18.8133 | -35.0491 | -28.9685
18.8194 | -35.0495 | -28.9758
18.8255 | -35.05 | -28.9831
18.8316 | -35.0504 | -28.9902
18.8377 | -35.0509 | -28.9974
18.8439 | -35.0513 | -29.0047
18.85 | -35.0517 | -29.012
18.8561 | -35.0521 | -29.0192
18.8622 | -35.0525 | -29.0265
18.8683 | -35.053 | -29.0335
18.8744 | -35.0534 | -29.0409
18.8805 | -35.0538 | -29.0482
18.8866 | -35.0543 | -29.0558
18.8927 | -35.0547 | -29.0628
18.8988 | -35.055 | -29.0698
18.9049 | -35.0554 | -29.0769
18.911 | -35.0559 | -29.0844
18.9171 | -35.0563 | -29.0916
18.9232 | -35.0567 | -29.0984
18.9293 | -35.0572 | -29.1058
18.9354 | -35.0576 | -29.1132
18.9415 | -35.058 | -29.1206
18.9476 | -35.0585 | -29.1279
18.9537 | -35.0589 | -29.135
18.9598 | -35.0593 | -29.142
18.9659 | -35.0597 | -29.149
18.972 | -35.0601 | -29.1564
18.9781 | -35.0605 | -29.1638
18.9843 | -35.0609 | -29.1714
18.9904 | -35.0614 | -29.1785
18.9965 | -35.0619 | -29.1854
19.0026 | -35.0624 | -29.1925
19.0087 | -35.0628 | -29.1999
19.0148 | -35.0632 | -29.2071
19.0209 | -35.0637 | -29.2144
19.027 | -35.0641 | -29.222
19.0331 | -35.0646 | -29.2292
19.0392 | -35.0649 | -29.2365
19.0453 | -35.0654 | -29.2437
19.0514 | -35.0658 | -29.251
19.0575 | -35.0663 | -29.2578
19.0636 | -35.0667 | -29.2651
19.0697 | -35.0671 | -29.2721
19.0758 | -35.0676 | -29.2794
19.0819 | -35.068 | -29.2867
19.088 | -35.0685 | -29.2937
19.0941 | -35.0689 | -29.301
19.1002 | -35.0693 | -29.308
19.1063 | -35.0697 | -29.3156
19.1124 | -35.0702 | -29.3229
19.1185 | -35.0706 | -29.3303
19.1246 | -35.071 | -29.3378
19.1308 | -35.0715 | -29.345
19.1369 | -35.072 | -29.3519
19.143 | -35.0724 | -29.3593
19.1491 | -35.0728 | -29.367
19.1552 | -35.0733 | -29.3742
19.1613 | -35.0737 | -29.3814
19.1674 | -35.0742 | -29.3887
19.1735 | -35.0746 | -29.396
19.1796 | -35.075 | -29.4028
19.1857 | -35.0755 | -29.4099
19.1918 | -35.0759 | -29.4167
19.1979 | -35.0763 | -29.424
19.204 | -35.0769 | -29.4311
19.2101 | -35.0772 | -29.4384
19.2162 | -35.0776 | -29.4457
19.2223 | -35.0781 | -29.4526
19.2284 | -35.0785 | -29.4596
19.2345 | -35.0789 | -29.4668
19.2406 | -35.0793 | -29.4741
19.2467 | -35.0798 | -29.4814
19.2528 | -35.0803 | -29.4891
19.2589 | -35.0807 | -29.4966
19.265 | -35.0812 | -29.5037
19.2712 | -35.0816 | -29.5107
19.2773 | -35.0821 | -29.5181
19.2834 | -35.0825 | -29.5251
19.2895 | -35.0829 | -29.5327
19.2956 | -35.0833 | -29.5402
19.3017 | -35.0838 | -29.5473
19.3078 | -35.0842 | -29.5545
19.3139 | -35.0846 | -29.5614
19.32 | -35.0851 | -29.5686
19.3261 | -35.0854 | -29.5759
19.3322 | -35.0859 | -29.5835
19.3383 | -35.0862 | -29.5907
19.3444 | -35.0867 | -29.5976
19.3505 | -35.0871 | -29.6049
19.3566 | -35.0876 | -29.6121
19.3627 | -35.0881 | -29.6188
19.3688 | -35.0885 | -29.6262
19.3749 | -35.089 | -29.6334
19.381 | -35.0894 | -29.6401
19.3871 | -35.0898 | -29.6472
19.3932 | -35.0903 | -29.6543
19.3993 | -35.0907 | -29.6616
19.4054 | -35.0911 | -29.6685
19.4115 | -35.0915 | -29.6755
19.4177 | -35.092 | -29.6822
19.4238 | -35.0924 | -29.6895
19.4299 | -35.0929 | -29.6968
19.436 | -35.0933 | -29.704
19.4421 | -35.0937 | -29.7113
19.4482 | -35.0941 | -29.7185
19.4543 | -35.0945 | -29.7253
19.4604 | -35.095 | -29.7325
19.4665 | -35.0954 | -29.7394
19.4726 | -35.0958 | -29.7464
19.4787 | -35.0963 | -29.7538
19.4848 | -35.0967 | -29.7612
19.4909 | -35.0971 | -29.7679
19.497 | -35.0975 | -29.7749
19.5031 | -35.0979 | -29.7824
19.5092 | -35.0984 | -29.7895
19.5153 | -35.0988 | -29.7962
19.5214 | -35.0992 | -29.8033
19.5275 | -35.0997 | -29.8107
19.5336 | -35.1002 | -29.8178
19.5397 | -35.1005 | -29.8248
19.5458 | -35.1009 | -29.8319
19.5519 | -35.1014 | -29.8387
19.5581 | -35.1019 | -29.8451
19.5642 | -35.1023 | -29.8529
19.5703 | -35.1027 | -29.86
19.5764 | -35.1032 | -29.867
19.5825 | -35.1036 | -29.8741
19.5886 | -35.104 | -29.8812
19.5947 | -35.1044 | -29.8883
19.6008 | -35.1048 | -29.8953
19.6069 | -35.1053 | -29.9025
19.613 | -35.1057 | -29.9098
19.6191 | -35.1062 | -29.9169
19.6252 | -35.1066 | -29.9241
19.6313 | -35.1071 | -29.9313
19.6374 | -35.1075 | -29.9386
19.6435 | -35.1079 | -29.9461
19.6496 | -35.1084 | -29.9535
19.6557 | -35.1088 | -29.9604
19.6618 | -35.1092 | -29.9672
19.6679 | -35.1096 | -29.974
19.674 | -35.1101 | -29.9814
19.6801 | -35.1105 | -29.9883
19.6862 | -35.111 | -29.9955
19.6923 | -35.1114 | -30.0026
19.6984 | -35.1118 | -30.0096
19.7046 | -35.1123 | -30.0166
19.7107 | -35.1127 | -30.024
19.7168 | -35.1131 | -30.0308
19.7229 | -35.1135 | -30.0378
19.729 | -35.114 | -30.0452
19.7351 | -35.1145 | -30.0521
19.7412 | -35.115 | -30.0592
19.7473 | -35.1154 | -30.0664
19.7534 | -35.1158 | -30.0732
19.7595 | -35.1162 | -30.0806
19.7656 | -35.1167 | -30.0876
19.7717 | -35.1171 | -30.0945
19.7778 | -35.1176 | -30.1016
19.7839 | -35.118 | -30.1087
19.79 | -35.1184 | -30.1159
19.7961 | -35.1188 | -30.123
19.8022 | -35.1193 | -30.1298
19.8083 | -35.1198 | -30.1369
19.8144 | -35.1202 | -30.144
19.8205 | -35.1207 | -30.1511
19.8266 | -35.1211 | -30.1579
19.8327 | -35.1214 | -30.1652
19.8388 | -35.1219 | -30.1723
19.845 | -35.1224 | -30.1794
19.8511 | -35.1228 | -30.1863
19.8572 | -35.1232 | -30.1933
19.8633 | -35.1237 | -30.2003
19.8694 | -35.1241 | -30.2071
19.8755 | -35.1246 | -30.2142
19.8816 | -35.1251 | -30.221
19.8877 | -35.1256 | -30.2283
19.8938 | -35.126 | -30.2352
19.8999 | -35.1264 | -30.2422
19.906 | -35.1269 | -30.2492
19.9121 | -35.1274 | -30.2562
19.9182 | -35.1278 | -30.2634
19.9243 | -35.1283 | -30.2704
19.9304 | -35.1287 | -30.2777
19.9365 | -35.1291 | -30.2846
19.9426 | -35.1296 | -30.2916
19.9487 | -35.1301 | -30.2986
19.9548 | -35.1306 | -30.3058
19.9609 | -35.131 | -30.3129
19.967 | -35.1315 | -30.3201
19.9731 | -35.132 | -30.3272
19.9792 | -35.1325 | -30.3347
19.9853 | -35.1329 | -30.3417
19.9915 | -35.1333 | -30.3486
19.9976 | -35.1338 | -30.3552
20.0037 | -35.1342 | -30.3625
20.0098 | -35.1347 | -30.3696
20.0159 | -35.1351 | -30.3761
20.022 | -35.1356 | -30.3833
20.0281 | -35.136 | -30.3906
20.0342 | -35.1365 | -30.3975
20.0403 | -35.1369 | -30.4046
20.0464 | -35.1374 | -30.4115
20.0525 | -35.1379 | -30.4185
20.0586 | -35.1384 | -30.4254
20.0647 | -35.1388 | -30.4324
20.0708 | -35.1393 | -30.4392
20.0769 | -35.1397 | -30.4462
20.083 | -35.1401 | -30.4532
20.0891 | -35.1406 | -30.4603
20.0952 | -35.1411 | -30.4672
20.1013 | -35.1415 | -30.4739
20.1074 | -35.1419 | -30.4811
20.1135 | -35.1424 | -30.4882
20.1196 | -35.1429 | -30.4946
20.1257 | -35.1434 | -30.5009
20.1319 | -35.1438 | -30.5077
20.138 | -35.1442 | -30.514
20.1441 | -35.1447 | -30.5208
20.1502 | -35.1451 | -30.5279
20.1563 | -35.1455 | -30.535
20.1624 | -35.1459 | -30.5419
20.1685 | -35.1464 | -30.5486
20.1746 | -35.1468 | -30.5557
20.1807 | -35.1473 | -30.5627
20.1868 | -35.1477 | -30.5699
20.1929 | -35.1481 | -30.5767
20.199 | -35.1485 | -30.5837
20.2051 | -35.149 | -30.5905
20.2112 | -35.1494 | -30.5976
20.2173 | -35.1498 | -30.6047
20.2234 | -35.1502 | -30.6116
20.2295 | -35.1507 | -30.6186
20.2356 | -35.1512 | -30.6258
20.2417 | -35.1516 | -30.6327
20.2478 | -35.152 | -30.6399
20.2539 | -35.1525 | -30.6471
20.26 | -35.153 | -30.654
20.2661 | -35.1535 | -30.6608
20.2723 | -35.1539 | -30.6678
20.2784 | -35.1543 | -30.6749
20.2845 | -35.1548 | -30.682
20.2906 | -35.1553 | -30.6884
20.2967 | -35.156 | -30.6949
20.3028 | -35.1564 | -30.7018
20.3089 | -35.1569 | -30.7087
20.315 | -35.1574 | -30.7158
20.3211 | -35.1578 | -30.7226
20.3272 | -35.1582 | -30.7293
20.3333 | -35.1586 | -30.7355
20.3394 | -35.1591 | -30.742
20.3455 | -35.1595 | -30.749
20.3516 | -35.1599 | -30.7564
20.3577 | -35.1603 | -30.7636
20.3638 | -35.1607 | -30.7702
20.3699 | -35.1612 | -30.7769
20.376 | -35.1617 | -30.7835
20.3821 | -35.1621 | -30.7902
20.3882 | -35.1625 | -30.7967
20.3943 | -35.163 | -30.8039
20.4004 | -35.1634 | -30.8108
20.4065 | -35.1638 | -30.8175
20.4126 | -35.1642 | -30.8242
20.4188 | -35.1647 | -30.8312
20.4249 | -35.1651 | -30.8379
20.431 | -35.1656 | -30.8448
20.4371 | -35.1659 | -30.8524
20.4432 | -35.1664 | -30.8593
20.4493 | -35.1669 | -30.8669
20.4554 | -35.1674 | -30.8734
20.4615 | -35.1678 | -30.88
20.4676 | -35.1682 | -30.8871
20.4737 | -35.1687 | -30.8941
20.4798 | -35.1691 | -30.9012
20.4859 | -35.1696 | -30.9082
20.492 | -35.17 | -30.915
20.4981 | -35.1704 | -30.9215
20.5042 | -35.1709 | -30.9282
20.5103 | -35.1714 | -30.9345
20.5164 | -35.1718 | -30.9411
20.5225 | -35.1722 | -30.948
20.5286 | -35.1727 | -30.9554
20.5347 | -35.1732 | -30.9625
20.5408 | -35.1736 | -30.9685
20.5469 | -35.1741 | -30.9754
20.553 | -35.1746 | -30.9821
20.5592 | -35.175 | -30.989
20.5653 | -35.1754 | -30.9959
20.5714 | -35.1759 | -31.0028
20.5775 | -35.1764 | -31.009
20.5836 | -35.1768 | -31.0159
20.5897 | -35.1773 | -31.0219
20.5958 | -35.1776 | -31.0281
20.6019 | -35.1781 | -31.034
20.608 | -35.1786 | -31.0415
20.6141 | -35.179 | -31.0478
20.6202 | -35.1794 | -31.0549
20.6263 | -35.1799 | -31.0615
20.6324 | -35.1802 | -31.068
20.6385 | -35.1807 | -31.0748
20.6446 | -35.1812 | -31.0818
20.6507 | -35.1817 | -31.0884
20.6568 | -35.1823 | -31.0953
20.6629 | -35.1827 | -31.1019
20.669 | -35.1831 | -31.1083
20.6751 | -35.1836 | -31.1155
20.6812 | -35.1841 | -31.1226
20.6873 | -35.1845 | -31.1296
20.6934 | -35.1849 | -31.1366
20.6995 | -35.1853 | -31.1429
20.7057 | -35.1859 | -31.1493
20.7118 | -35.1863 | -31.156
20.7179 | -35.1866 | -31.162
20.724 | -35.1871 | -31.1688
20.7301 | -35.1875 | -31.1759
20.7362 | -35.188 | -31.1832
20.7423 | -35.1884 | -31.1897
20.7484 | -35.1889 | -31.1965
20.7545 | -35.1891 | -31.2032
20.7606 | -35.1897 | -31.2097
20.7667 | -35.1901 | -31.2162
20.7728 | -35.1905 | -31.2227
20.7789 | -35.191 | -31.2296
20.785 | -35.1914 | -31.2361
20.7911 | -35.1919 | -31.2426
20.7972 | -35.1923 | -31.2497
20.8033 | -35.1928 | -31.2568
20.8094 | -35.1933 | -31.2638
20.8155 | -35.1938 | -31.27
20.8216 | -35.1942 | -31.277
20.8277 | -35.1947 | -31.2829
20.8338 | -35.1952 | -31.2893
20.8399 | -35.1956 | -31.296
20.8461 | -35.1961 | -31.3027
20.8522 | -35.1965 | -31.3095
20.8583 | -35.1969 | -31.316
20.8644 | -35.1974 | -31.3225
20.8705 | -35.1979 | -31.3295
20.8766 | -35.1982 | -31.3367
20.8827 | -35.1987 | -31.3434
20.8888 | -35.1991 | -31.3502
20.8949 | -35.1995 | -31.3573
20.901 | -35.2 | -31.3631
20.9071 | -35.2003 | -31.3703
20.9132 | -35.2007 | -31.3762
20.9193 | -35.2011 | -31.3833
20.9254 | -35.2016 | -31.3902
20.9315 | -35.202 | -31.3956
20.9376 | -35.2024 | -31.4022
20.9437 | -35.2028 | -31.4089
20.9498 | -35.2033 | -31.4153
20.9559 | -35.2038 | -31.422
20.962 | -35.2043 | -31.4287
20.9681 | -35.2048 | -31.4351
20.9742 | -35.2052 | -31.4413
20.9803 | -35.2057 | -31.4477
20.9864 | -35.2063 | -31.4544
20.9926 | -35.2068 | -31.4613
20.9987 | -35.2073 | -31.4683
21.0048 | -35.2078 | -31.4746
21.0109 | -35.2082 | -31.4814
21.017 | -35.2086 | -31.4884
21.0231 | -35.209 | -31.4948
21.0292 | -35.2095 | -31.5014
21.0353 | -35.21 | -31.5084
21.0414 | -35.2104 | -31.5151
21.0475 | -35.2108 | -31.5219
21.0536 | -35.2112 | -31.5291
21.0597 | -35.2115 | -31.5365
21.0658 | -35.2119 | -31.5443
21.0719 | -35.2125 | -31.5512
21.078 | -35.2128 | -31.5579
21.0841 | -35.2132 | -31.5648
21.0902 | -35.2138 | -31.5712
21.0963 | -35.214 | -31.5779
21.1024 | -35.2144 | -31.5844
21.1085 | -35.2149 | -31.5911
21.1146 | -35.2154 | -31.5984
21.1207 | -35.2158 | -31.6052
21.1268 | -35.2163 | -31.6122
21.133 | -35.2167 | -31.619
21.1391 | -35.2172 | -31.6257
21.1452 | -35.2177 | -31.6328
21.1513 | -35.2182 | -31.6394
21.1574 | -35.2187 | -31.6462
21.1635 | -35.2192 | -31.6524
21.1696 | -35.2197 | -31.659
21.1757 | -35.2201 | -31.6658
21.1818 | -35.2207 | -31.6726
21.1879 | -35.2211 | -31.6792
21.194 | -35.2216 | -31.6855
21.2001 | -35.222 | -31.6927
21.2062 | -35.2225 | -31.6995
21.2123 | -35.2232 | -31.7065
21.2184 | -35.2235 | -31.7134
21.2245 | -35.2239 | -31.7194
21.2306 | -35.2243 | -31.7259
21.2367 | -35.2248 | -31.7328
21.2428 | -35.2253 | -31.7393
21.2489 | -35.2257 | -31.7459
21.255 | -35.2261 | -31.7524
21.2611 | -35.2266 | -31.7592
21.2672 | -35.227 | -31.766
21.2733 | -35.2273 | -31.7727
21.2795 | -35.2277 | -31.7787
21.2856 | -35.2282 | -31.7853
21.2917 | -35.2286 | -31.7916
21.2978 | -35.229 | -31.7983
21.3039 | -35.2295 | -31.8054
21.31 | -35.23 | -31.8118
21.3161 | -35.2304 | -31.8187
21.3222 | -35.2309 | -31.8254
21.3283 | -35.2314 | -31.8321
21.3344 | -35.2318 | -31.8389
21.3405 | -35.2323 | -31.8456
21.3466 | -35.2327 | -31.8521
21.3527 | -35.2332 | -31.8591
21.3588 | -35.2337 | -31.8658
21.3649 | -35.2341 | -31.8726
21.371 | -35.2344 | -31.8786
21.3771 | -35.2348 | -31.8846
21.3832 | -35.2352 | -31.892
21.3893 | -35.2355 | -31.8979
21.3954 | -35.236 | -31.9053
21.4015 | -35.2365 | -31.9116
21.4076 | -35.237 | -31.9184
21.4137 | -35.2375 | -31.9252
21.4199 | -35.238 | -31.9314
21.426 | -35.2385 | -31.9377
21.4321 | -35.239 | -31.944
21.4382 | -35.2394 | -31.9503
21.4443 | -35.2399 | -31.9563
21.4504 | -35.2405 | -31.9629
21.4565 | -35.2408 | -31.9686
21.4626 | -35.2413 | -31.9756
21.4687 | -35.2417 | -31.982
21.4748 | -35.2421 | -31.9883
21.4809 | -35.2425 | -31.995
21.487 | -35.243 | -32.002
21.4931 | -35.2435 | -32.0077
21.4992 | -35.244 | -32.0146
21.5053 | -35.2444 | -32.0216
21.5114 | -35.245 | -32.0275
21.5175 | -35.2454 | -32.0341
21.5236 | -35.246 | -32.0405
21.5297 | -35.2465 | -32.0474
21.5358 | -35.247 | -32.0549
21.5419 | -35.2474 | -32.0614
21.548 | -35.2478 | -32.0671
21.5541 | -35.2482 | -32.0739
21.5602 | -35.2487 | -32.0809
21.5664 | -35.2492 | -32.0869
21.5725 | -35.2496 | -32.093
21.5786 | -35.25 | -32.0995
21.5847 | -35.2505 | -32.1061
21.5908 | -35.251 | -32.1127
21.5969 | -35.2515 | -32.1187
21.603 | -35.2519 | -32.1246
21.6091 | -35.2524 | -32.131
21.6152 | -35.2528 | -32.1377
21.6213 | -35.2533 | -32.1441
21.6274 | -35.2537 | -32.1501
21.6335 | -35.2543 | -32.1563
21.6396 | -35.2547 | -32.1632
21.6457 | -35.2552 | -32.1705
21.6518 | -35.2556 | -32.1769
21.6579 | -35.2561 | -32.1833
21.664 | -35.2565 | -32.1895
21.6701 | -35.257 | -32.1957
21.6762 | -35.2573 | -32.2019
21.6823 | -35.2579 | -32.2075
21.6884 | -35.2585 | -32.2137
21.6945 | -35.2589 | -32.2201
21.7006 | -35.2594 | -32.2267
21.7068 | -35.2598 | -32.2329
21.7129 | -35.2603 | -32.2392
21.719 | -35.2607 | -32.246
21.7251 | -35.2613 | -32.2528
21.7312 | -35.2617 | -32.2596
21.7373 | -35.2622 | -32.2656
21.7434 | -35.2626 | -32.2717
21.7495 | -35.2629 | -32.2772
21.7556 | -35.263 | -32.2841
21.7617 | -35.2649 | -32.2899
21.7678 | -35.2651 | -32.2958
21.7739 | -35.2656 | -32.3016
21.78 | -35.2661 | -32.3075
21.7861 | -35.2667 | -32.3138
21.7922 | -35.2672 | -32.3204
21.7983 | -35.2675 | -32.327
21.8044 | -35.268 | -32.3328
21.8105 | -35.2686 | -32.3392
21.8166 | -35.2691 | -32.3457
21.8227 | -35.2696 | -32.3518
21.8288 | -35.2702 | -32.3577
21.8349 | -35.2708 | -32.3635
21.841 | -35.2712 | -32.3704
21.8471 | -35.2718 | -32.3745
21.8533 | -35.2722 | -32.3804
21.8594 | -35.2725 | -32.3872
21.8655 | -35.273 | -32.393
21.8716 | -35.2735 | -32.3985
21.8777 | -35.274 | -32.4048
21.8838 | -35.2745 | -32.4106
21.8899 | -35.2749 | -32.4167
21.896 | -35.2752 | -32.4232
21.9021 | -35.2757 | -32.4297
21.9082 | -35.2761 | -32.4356
21.9143 | -35.2766 | -32.4411
21.9204 | -35.277 | -32.4471
21.9265 | -35.2776 | -32.4534
21.9326 | -35.278 | -32.4602
21.9387 | -35.2786 | -32.4666
21.9448 | -35.2791 | -32.4733
21.9509 | -35.2797 | -32.4796
21.957 | -35.2801 | -32.4857
21.9631 | -35.2807 | -32.4918
21.9692 | -35.2813 | -32.4985
21.9753 | -35.2817 | -32.5044
21.9814 | -35.2822 | -32.511
21.9875 | -35.2826 | -32.5198
21.9937 | -35.2832 | -32.5273
21.9998 | -35.2838 | -32.5344
22.0059 | -35.2842 | -32.5409
22.012 | -35.2846 | -32.5478
22.0181 | -35.2851 | -32.5544
22.0242 | -35.2856 | -32.5615
22.0303 | -35.286 | -32.5675
22.0364 | -35.2866 | -32.5733
22.0425 | -35.2871 | -32.5802
22.0486 | -35.2875 | -32.5862
22.0547 | -35.2879 | -32.5928
22.0608 | -35.2884 | -32.599
22.0669 | -35.2888 | -32.6055
22.073 | -35.2892 | -32.6116
22.0791 | -35.2897 | -32.6182
22.0852 | -35.2901 | -32.6244
22.0913 | -35.2906 | -32.6306
22.0974 | -35.2911 | -32.6366
22.1035 | -35.2915 | -32.6429
22.1096 | -35.292 | -32.6498
22.1157 | -35.2925 | -32.6565
22.1218 | -35.2929 | -32.6628
22.1279 | -35.2933 | -32.6688
22.134 | -35.2938 | -32.6747
22.1402 | -35.2943 | -32.6817
22.1463 | -35.2947 | -32.6895
22.1524 | -35.2951 | -32.6947
22.1585 | -35.2956 | -32.7013
22.1646 | -35.2961 | -32.7065
22.1707 | -35.2966 | -32.7134
22.1768 | -35.2972 | -32.7193
22.1829 | -35.2976 | -32.7263
22.189 | -35.2981 | -32.7326
22.1951 | -35.2985 | -32.7379
22.2012 | -35.299 | -32.7442
22.2073 | -35.2995 | -32.7504
22.2134 | -35.2999 | -32.7567
22.2195 | -35.3004 | -32.7624
22.2256 | -35.3009 | -32.7689
22.2317 | -35.3013 | -32.7746
22.2378 | -35.3018 | -32.7813
22.2439 | -35.3023 | -32.7879
22.25 | -35.3027 | -32.7942
22.2561 | -35.3033 | -32.8009
22.2622 | -35.3037 | -32.8064
22.2683 | -35.3043 | -32.8122
22.2744 | -35.3048 | -32.8189
22.2806 | -35.3052 | -32.8247
22.2867 | -35.3057 | -32.8313
22.2928 | -35.3062 | -32.8385
22.2989 | -35.3066 | -32.8437
22.305 | -35.3074 | -32.8483
22.3111 | -35.308 | -32.8549
22.3172 | -35.3085 | -32.8617
22.3233 | -35.309 | -32.868
22.3294 | -35.3095 | -32.8751
22.3355 | -35.3099 | -32.8812
22.3416 | -35.3103 | -32.8875
22.3477 | -35.3108 | -32.8935
22.3538 | -35.3113 | -32.8997
22.3599 | -35.3117 | -32.9062
22.366 | -35.3123 | -32.9132
22.3721 | -35.3127 | -32.9185
22.3782 | -35.3131 | -32.9247
22.3843 | -35.3135 | -32.9312
22.3904 | -35.314 | -32.9377
22.3965 | -35.3146 | -32.9444
22.4026 | -35.315 | -32.9506
22.4087 | -35.3156 | -32.9559
22.4148 | -35.3161 | -32.962
22.4209 | -35.3166 | -32.9679
22.4271 | -35.317 | -32.9744
22.4332 | -35.3174 | -32.9809
22.4393 | -35.3179 | -32.9864
22.4454 | -35.3185 | -32.9926
22.4515 | -35.3191 | -32.9989
22.4576 | -35.3196 | -33.0036
22.4637 | -35.32 | -33.0115
22.4698 | -35.3207 | -33.0173
22.4759 | -35.3213 | -33.0222
22.482 | -35.3217 | -33.0283
22.4881 | -35.3222 | -33.0343
22.4942 | -35.3227 | -33.0407
22.5003 | -35.3231 | -33.0473
22.5064 | -35.3236 | -33.0536
22.5125 | -35.324 | -33.0591
22.5186 | -35.3246 | -33.0651
22.5247 | -35.3251 | -33.0714
22.5308 | -35.3256 | -33.078
22.5369 | -35.3261 | -33.084
22.543 | -35.3266 | -33.0908
22.5491 | -35.3271 | -33.0962
22.5552 | -35.3277 | -33.1024
22.5613 | -35.3281 | -33.1088
22.5675 | -35.3286 | -33.1151
22.5736 | -35.3291 | -33.1214
22.5797 | -35.3295 | -33.1273
22.5858 | -35.33 | -33.1338
22.5919 | -35.3304 | -33.1395
22.598 | -35.331 | -33.1457
22.6041 | -35.3316 | -33.1518
22.6102 | -35.3321 | -33.1568
22.6163 | -35.3324 | -33.1632
22.6224 | -35.3329 | -33.1689
22.6285 | -35.3335 | -33.1742
22.6346 | -35.3338 | -33.1802
22.6407 | -35.3343 | -33.1861
22.6468 | -35.3349 | -33.1923
22.6529 | -35.3353 | -33.1973
22.659 | -35.3358 | -33.2036
22.6651 | -35.3363 | -33.2093
22.6712 | -35.3367 | -33.216
22.6773 | -35.3372 | -33.2223
22.6834 | -35.3376 | -33.2284
22.6895 | -35.3382 | -33.2333
22.6956 | -35.3387 | -33.2398
22.7017 | -35.3391 | -33.246
22.7079 | -35.3395 | -33.2521
22.714 | -35.3401 | -33.2585
22.7201 | -35.3406 | -33.2647
22.7262 | -35.3413 | -33.2704
22.7323 | -35.3418 | -33.2764
22.7384 | -35.3423 | -33.2828
22.7445 | -35.3429 | -33.2887
22.7506 | -35.3434 | -33.2952
22.7567 | -35.3438 | -33.3013
22.7628 | -35.3442 | -33.3073
22.7689 | -35.3446 | -33.3138
22.775 | -35.3452 | -33.3191
22.7811 | -35.3456 | -33.3252
22.7872 | -35.3462 | -33.3311
22.7933 | -35.3468 | -33.3362
22.7994 | -35.3473 | -33.3426
22.8055 | -35.3477 | -33.3491
22.8116 | -35.3482 | -33.3549
22.8177 | -35.3488 | -33.3609
22.8238 | -35.3492 | -33.3669
22.8299 | -35.3498 | -33.3726
22.836 | -35.3502 | -33.379
22.8421 | -35.3508 | -33.3854
22.8482 | -35.3513 | -33.3911
22.8544 | -35.3518 | -33.3976
22.8605 | -35.3523 | -33.4039
22.8666 | -35.3528 | -33.4095
22.8727 | -35.3533 | -33.4153
22.8788 | -35.3538 | -33.4211
22.8849 | -35.3543 | -33.4273
22.891 | -35.3548 | -33.4333
22.8971 | -35.3553 | -33.4389
22.9032 | -35.3557 | -33.4452
22.9093 | -35.3562 | -33.4513
22.9154 | -35.3566 | -33.4575
22.9215 | -35.3571 | -33.4644
22.9276 | -35.3575 | -33.4714
22.9337 | -35.358 | -33.4761
22.9398 | -35.3585 | -33.4823
22.9459 | -35.359 | -33.4886
22.952 | -35.3595 | -33.4947
22.9581 | -35.36 | -33.5004
22.9642 | -35.3605 | -33.5062
22.9703 | -35.3611 | -33.5117
22.9764 | -35.3617 | -33.5177
22.9825 | -35.3621 | -33.5236
22.9886 | -35.3625 | -33.5294
22.9948 | -35.363 | -33.5355
23.0009 | -35.3635 | -33.5415
23.007 | -35.364 | -33.547
23.0131 | -35.3645 | -33.553
23.0192 | -35.365 | -33.5592
23.0253 | -35.3655 | -33.5656
23.0314 | -35.366 | -33.5717
23.0375 | -35.3666 | -33.5773
23.0436 | -35.3671 | -33.5832
23.0497 | -35.3676 | -33.5892
23.0558 | -35.3681 | -33.5952
23.0619 | -35.3686 | -33.6008
23.068 | -35.3691 | -33.6067
23.0741 | -35.3696 | -33.6133
23.0802 | -35.37 | -33.6187
23.0863 | -35.3705 | -33.6246
23.0924 | -35.371 | -33.6301
23.0985 | -35.3715 | -33.6364
23.1046 | -35.372 | -33.6424
23.1107 | -35.3725 | -33.6481
23.1168 | -35.373 | -33.6539
23.1229 | -35.3735 | -33.6597
23.129 | -35.3741 | -33.6659
23.1351 | -35.3746 | -33.6722
23.1413 | -35.375 | -33.6785
23.1474 | -35.3755 | -33.6842
23.1535 | -35.376 | -33.6906
23.1596 | -35.3765 | -33.6968
23.1657 | -35.377 | -33.7029
23.1718 | -35.3776 | -33.7089
23.1779 | -35.3781 | -33.7154
23.184 | -35.3785 | -33.7207
23.1901 | -35.3791 | -33.7265
23.1962 | -35.3795 | -33.7323
23.2023 | -35.38 | -33.7386
23.2084 | -35.3805 | -33.7446
23.2145 | -35.3811 | -33.7503
23.2206 | -35.3815 | -33.7563
23.2267 | -35.382 | -33.7621
23.2328 | -35.3825 | -33.7675
23.2389 | -35.3829 | -33.773
23.245 | -35.3835 | -33.7792
23.2511 | -35.384 | -33.7851
23.2572 | -35.3846 | -33.7911
23.2633 | -35.3851 | -33.7972
23.2694 | -35.3855 | -33.8027
23.2755 | -35.3861 | -33.8087
23.2817 | -35.3866 | -33.8144
23.2878 | -35.3871 | -33.8205
23.2939 | -35.3876 | -33.8268
23.3 | -35.3882 | -33.8331
23.3061 | -35.3887 | -33.8387
23.3122 | -35.3893 | -33.8448
23.3183 | -35.3898 | -33.851
23.3244 | -35.3903 | -33.8569
23.3305 | -35.3908 | -33.8629
23.3366 | -35.3913 | -33.8683
23.3427 | -35.3918 | -33.8742
23.3488 | -35.3923 | -33.8797
23.3549 | -35.3927 | -33.8861
23.361 | -35.3933 | -33.8915
23.3671 | -35.3939 | -33.8973
23.3732 | -35.3943 | -33.9031
23.3793 | -35.3949 | -33.9094
23.3854 | -35.3955 | -33.9152
23.3915 | -35.3961 | -33.9205
23.3976 | -35.3966 | -33.9275
23.4037 | -35.3971 | -33.9332
23.4098 | -35.3977 | -33.9391
23.4159 | -35.3982 | -33.9447
23.422 | -35.3987 | -33.9507
23.4282 | -35.3992 | -33.9563
23.4343 | -35.3996 | -33.9621
23.4404 | -35.4001 | -33.968
23.4465 | -35.4006 | -33.9734
23.4526 | -35.4011 | -33.9797
23.4587 | -35.4017 | -33.985
23.4648 | -35.4021 | -33.9906
23.4709 | -35.4026 | -33.9969
23.477 | -35.4031 | -34.0029
23.4831 | -35.4037 | -34.0087
23.4892 | -35.4042 | -34.0149
23.4953 | -35.4048 | -34.021
23.5014 | -35.4053 | -34.0269
23.5075 | -35.4058 | -34.0327
23.5136 | -35.4063 | -34.0389
23.5197 | -35.4068 | -34.0448
23.5258 | -35.4072 | -34.0506
23.5319 | -35.4077 | -34.0567
23.538 | -35.4082 | -34.0628
23.5441 | -35.4087 | -34.0689
23.5502 | -35.4093 | -34.0749
23.5563 | -35.4099 | -34.0808
23.5624 | -35.4104 | -34.0865
23.5686 | -35.411 | -34.0921
23.5747 | -35.4115 | -34.0976
23.5808 | -35.4121 | -34.1031
23.5869 | -35.4125 | -34.1087
23.593 | -35.4131 | -34.1144
23.5991 | -35.4136 | -34.1199
23.6052 | -35.4142 | -34.1253
23.6113 | -35.4147 | -34.1314
23.6174 | -35.4152 | -34.1374
23.6235 | -35.4157 | -34.1429
23.6296 | -35.4161 | -34.1486
23.6357 | -35.4166 | -34.154
23.6418 | -35.4172 | -34.1599
23.6479 | -35.4176 | -34.1655
23.654 | -35.4181 | -34.1713
23.6601 | -35.4186 | -34.177
23.6662 | -35.4191 | -34.1826
23.6723 | -35.4197 | -34.1878
23.6784 | -35.4203 | -34.1936
23.6845 | -35.4208 | -34.1995
23.6906 | -35.4213 | -34.2051
23.6967 | -35.4217 | -34.2106
23.7028 | -35.4223 | -34.2163
23.7089 | -35.4228 | -34.2219
23.7151 | -35.4233 | -34.2275
23.7212 | -35.4237 | -34.233
23.7273 | -35.4243 | -34.2388
23.7334 | -35.4248 | -34.2449
23.7395 | -35.4253 | -34.2504
23.7456 | -35.4258 | -34.2561
23.7517 | -35.4263 | -34.2619
23.7578 | -35.4268 | -34.2672
23.7639 | -35.4273 | -34.2731
23.77 | -35.4278 | -34.2792
23.7761 | -35.4284 | -34.2853
23.7822 | -35.429 | -34.2905
23.7883 | -35.4295 | -34.2964
23.7944 | -35.43 | -34.302
23.8005 | -35.4306 | -34.3081
23.8066 | -35.4311 | -34.3136
23.8127 | -35.4317 | -34.3192
23.8188 | -35.4322 | -34.3251
23.8249 | -35.4327 | -34.3312
23.831 | -35.4333 | -34.3372
23.8371 | -35.4338 | -34.3429
23.8432 | -35.4343 | -34.3485
23.8493 | -35.4348 | -34.3534
23.8555 | -35.4352 | -34.3589
23.8616 | -35.4357 | -34.3651
23.8677 | -35.4362 | -34.3712
23.8738 | -35.4368 | -34.3768
23.8799 | -35.4373 | -34.3827
23.886 | -35.4377 | -34.3875
23.8921 | -35.4383 | -34.3935
23.8982 | -35.4388 | -34.3991
23.9043 | -35.4393 | -34.405
23.9104 | -35.4398 | -34.4108
23.9165 | -35.4403 | -34.4162
23.9226 | -35.4409 | -34.4219
23.9287 | -35.4414 | -34.4276
23.9348 | -35.4419 | -34.4335
23.9409 | -35.4424 | -34.4392
23.947 | -35.4429 | -34.4447
23.9531 | -35.4434 | -34.4497
23.9592 | -35.4439 | -34.4555
23.9653 | -35.4444 | -34.4611
23.9714 | -35.4449 | -34.4666
23.9775 | -35.4454 | -34.4722
23.9836 | -35.4459 | -34.4777
23.9897 | -35.4465 | -34.4832
23.9958 | -35.447 | -34.4884
24.002 | -35.4476 | -34.4944
24.0081 | -35.4481 | -34.5007
24.0142 | -35.4486 | -34.5064
24.0203 | -35.4491 | -34.5119
24.0264 | -35.4496 | -34.5175
24.0325 | -35.4501 | -34.523
24.0386 | -35.4506 | -34.5285
24.0447 | -35.4511 | -34.5346
24.0508 | -35.4517 | -34.5402
24.0569 | -35.4522 | -34.5461
24.063 | -35.4528 | -34.5511
24.0691 | -35.4532 | -34.5568
24.0752 | -35.4537 | -34.5626
24.0813 | -35.4542 | -34.5683
24.0874 | -35.4547 | -34.5742
24.0935 | -35.4552 | -34.58
24.0996 | -35.4557 | -34.5854
24.1057 | -35.4563 | -34.5906
24.1118 | -35.4569 | -34.5961
24.1179 | -35.4574 | -34.6017
24.124 | -35.458 | -34.6077
24.1301 | -35.4585 | -34.6134
24.1362 | -35.459 | -34.6191
24.1424 | -35.4595 | -34.6244
24.1485 | -35.4601 | -34.6303
24.1546 | -35.4605 | -34.6359
24.1607 | -35.461 | -34.6417
24.1668 | -35.4615 | -34.647
24.1729 | -35.462 | -34.6525
24.179 | -35.4625 | -34.6578
24.1851 | -35.4631 | -34.6637
24.1912 | -35.4635 | -34.6688
24.1973 | -35.464 | -34.6745
24.2034 | -35.4646 | -34.6796
24.2095 | -35.4651 | -34.6853
24.2156 | -35.4656 | -34.6908
24.2217 | -35.4661 | -34.6958
24.2278 | -35.4666 | -34.7011
24.2339 | -35.4671 | -34.7068
24.24 | -35.4676 | -34.7125
24.2461 | -35.4681 | -34.7186
24.2522 | -35.4686 | -34.7245
24.2583 | -35.4691 | -34.7305
24.2644 | -35.4697 | -34.7363
24.2705 | -35.4702 | -34.7417
24.2766 | -35.4708 | -34.7476
24.2827 | -35.4713 | -34.7526
24.2889 | -35.4719 | -34.7585
24.295 | -35.4724 | -34.7642
24.3011 | -35.4729 | -34.7697
24.3072 | -35.4734 | -34.7754
24.3133 | -35.4739 | -34.7806
24.3194 | -35.4744 | -34.7864
24.3255 | -35.475 | -34.7919
24.3316 | -35.4755 | -34.7973
24.3377 | -35.4759 | -34.8029
24.3438 | -35.4764 | -34.808
24.3499 | -35.4769 | -34.8134
24.356 | -35.4774 | -34.8191
24.3621 | -35.4779 | -34.8247
24.3682 | -35.4784 | -34.8305
24.3743 | -35.4789 | -34.8362
24.3804 | -35.4794 | -34.842
24.3865 | -35.4799 | -34.8471
24.3926 | -35.4804 | -34.853
24.3987 | -35.481 | -34.8588
24.4048 | -35.4815 | -34.8646
24.4109 | -35.482 | -34.8704
24.417 | -35.4825 | -34.876
24.4231 | -35.4829 | -34.8821
24.4293 | -35.4835 | -34.8877
24.4354 | -35.484 | -34.8932
24.4415 | -35.4846 | -34.8988
24.4476 | -35.4851 | -34.9044
24.4537 | -35.4856 | -34.9101
24.4598 | -35.4861 | -34.9153
24.4659 | -35.4867 | -34.9209
24.472 | -35.4871 | -34.9264
24.4781 | -35.4876 | -34.9316
24.4842 | -35.4881 | -34.9372
24.4903 | -35.4886 | -34.943
24.4964 | -35.4891 | -34.9487
24.5025 | -35.4896 | -34.954
24.5086 | -35.4901 | -34.9595
24.5147 | -35.4906 | -34.9652
24.5208 | -35.4911 | -34.9706
24.5269 | -35.4916 | -34.976
24.533 | -35.4921 | -34.9819
24.5391 | -35.4927 | -34.9874
24.5452 | -35.4931 | -34.9932
24.5513 | -35.4936 | -34.9989
24.5574 | -35.4942 | -35.004
24.5635 | -35.4947 | -35.0092
24.5696 | -35.4952 | -35.0149
24.5758 | -35.4957 | -35.0203
24.5819 | -35.4963 | -35.0258
24.588 | -35.4967 | -35.0311
24.5941 | -35.4972 | -35.0363
24.6002 | -35.4977 | -35.0417
24.6063 | -35.4982 | -35.047
24.6124 | -35.4987 | -35.0527
24.6185 | -35.4992 | -35.0581
24.6246 | -35.4997 | -35.0638
24.6307 | -35.5002 | -35.0693
24.6368 | -35.5007 | -35.0752
24.6429 | -35.5013 | -35.0806
24.649 | -35.5018 | -35.0862
24.6551 | -35.5023 | -35.0918
24.6612 | -35.5027 | -35.0978
24.6673 | -35.5031 | -35.1037
24.6734 | -35.5036 | -35.1087
24.6795 | -35.5042 | -35.1144
24.6856 | -35.5047 | -35.1195
24.6917 | -35.5051 | -35.1245
24.6978 | -35.5056 | -35.1294
24.7039 | -35.5061 | -35.1352
24.71 | -35.5066 | -35.1406
24.7162 | -35.5071 | -35.146
24.7223 | -35.5077 | -35.1515
24.7284 | -35.5082 | -35.1565
24.7345 | -35.5086 | -35.1623
24.7406 | -35.5091 | -35.1671
24.7467 | -35.5096 | -35.173
24.7528 | -35.5101 | -35.1782
24.7589 | -35.5107 | -35.1835
24.765 | -35.5112 | -35.1897
24.7711 | -35.5117 | -35.1948
24.7772 | -35.5123 | -35.2011
24.7833 | -35.5128 | -35.2069
24.7894 | -35.5134 | -35.2125
24.7955 | -35.5138 | -35.2178
24.8016 | -35.5142 | -35.223
24.8077 | -35.5147 | -35.2284
24.8138 | -35.5152 | -35.2339
24.8199 | -35.5157 | -35.2397
24.826 | -35.5161 | -35.2452
24.8321 | -35.5165 | -35.2506
24.8382 | -35.517 | -35.256
24.8443 | -35.5175 | -35.2616
24.8504 | -35.5181 | -35.267
24.8565 | -35.5186 | -35.2723
24.8627 | -35.5191 | -35.2777
24.8688 | -35.5195 | -35.2829
24.8749 | -35.52 | -35.2883
24.881 | -35.5205 | -35.2931
24.8871 | -35.5209 | -35.2989
24.8932 | -35.5214 | -35.3049
24.8993 | -35.5218 | -35.3101
24.9054 | -35.5222 | -35.3157
24.9115 | -35.5228 | -35.3211
24.9176 | -35.5232 | -35.3269
24.9237 | -35.5237 | -35.3322
24.9298 | -35.5243 | -35.3383
24.9359 | -35.5248 | -35.3443
24.942 | -35.5252 | -35.3498
24.9481 | -35.5257 | -35.3556
24.9542 | -35.5262 | -35.3612
24.9603 | -35.5266 | -35.3667
24.9664 | -35.5271 | -35.3723
24.9725 | -35.5276 | -35.378
24.9786 | -35.5281 | -35.3829
24.9847 | -35.5286 | -35.3883
24.9908 | -35.5291 | -35.3943
24.9969 | -35.5296 | -35.3997
25.0031 | -35.53 | -35.4053
25.0092 | -35.5306 | -35.4106
25.0153 | -35.531 | -35.4157
25.0214 | -35.5315 | -35.4216
25.0275 | -35.532 | -35.4268
25.0336 | -35.5325 | -35.4319
25.0397 | -35.5331 | -35.4372
25.0458 | -35.5336 | -35.4422
25.0519 | -35.5341 | -35.4479
25.058 | -35.5346 | -35.4532
25.0641 | -35.5351 | -35.4589
25.0702 | -35.5356 | -35.4642
25.0763 | -35.536 | -35.4696
25.0824 | -35.5365 | -35.4749
25.0885 | -35.537 | -35.4805
25.0946 | -35.5375 | -35.4858
25.1007 | -35.5379 | -35.4915
25.1068 | -35.5384 | -35.4972
25.1129 | -35.5389 | -35.5028
25.119 | -35.5393 | -35.508
25.1251 | -35.5397 | -35.5131
25.1312 | -35.5402 | -35.5189
25.1373 | -35.5407 | -35.524
25.1435 | -35.5412 | -35.5291
25.1496 | -35.5417 | -35.5341
25.1557 | -35.5422 | -35.5396
25.1618 | -35.5426 | -35.5453
25.1679 | -35.5431 | -35.5509
25.174 | -35.5437 | -35.5565
25.1801 | -35.5441 | -35.5625
25.1862 | -35.5446 | -35.5677
25.1923 | -35.545 | -35.5727
25.1984 | -35.5455 | -35.5782
25.2045 | -35.5459 | -35.584
25.2106 | -35.5464 | -35.5897
25.2167 | -35.5467 | -35.5949
25.2228 | -35.5473 | -35.6003
25.2289 | -35.5477 | -35.6058
25.235 | -35.5482 | -35.611
25.2411 | -35.5486 | -35.6164
25.2472 | -35.549 | -35.6219
25.2533 | -35.5494 | -35.6275
25.2594 | -35.5498 | -35.633
25.2655 | -35.5503 | -35.6382
25.2716 | -35.5508 | -35.6435
25.2777 | -35.5512 | -35.6489
25.2838 | -35.5517 | -35.6542
25.29 | -35.5522 | -35.6594
25.2961 | -35.5526 | -35.665
25.3022 | -35.5531 | -35.6705
25.3083 | -35.5535 | -35.676
25.3144 | -35.554 | -35.6813
25.3205 | -35.5544 | -35.6861
25.3266 | -35.5548 | -35.691
25.3327 | -35.5552 | -35.6967
25.3388 | -35.5557 | -35.7022
25.3449 | -35.5561 | -35.7076
25.351 | -35.5566 | -35.713
25.3571 | -35.5571 | -35.7184
25.3632 | -35.5575 | -35.723
25.3693 | -35.558 | -35.7282
25.3754 | -35.5584 | -35.7338
25.3815 | -35.5588 | -35.7393
25.3876 | -35.5592 | -35.7444
25.3937 | -35.5597 | -35.7498
25.3998 | -35.5601 | -35.7542
25.4059 | -35.5606 | -35.7596
25.412 | -35.561 | -35.7647
25.4181 | -35.5615 | -35.7699
25.4242 | -35.5619 | -35.7754
25.4304 | -35.5624 | -35.7806
25.4365 | -35.5628 | -35.786
25.4426 | -35.5633 | -35.7909
25.4487 | -35.5638 | -35.7963
25.4548 | -35.5643 | -35.8022
25.4609 | -35.5648 | -35.8073
25.467 | -35.5651 | -35.8123
25.4731 | -35.5656 | -35.8177
25.4792 | -35.5661 | -35.8234
25.4853 | -35.5666 | -35.829
25.4914 | -35.567 | -35.8344
25.4975 | -35.5675 | -35.84
25.5036 | -35.5679 | -35.8452
25.5097 | -35.5684 | -35.851
25.5158 | -35.5688 | -35.8559
25.5219 | -35.5693 | -35.8612
25.528 | -35.5697 | -35.8666
25.5341 | -35.5702 | -35.8723
25.5402 | -35.5706 | -35.8779
25.5463 | -35.5711 | -35.8832
25.5524 | -35.5715 | -35.8888
25.5585 | -35.572 | -35.8943
25.5646 | -35.5723 | -35.8998
25.5707 | -35.5728 | -35.905
25.5769 | -35.5731 | -35.9092
25.583 | -35.5735 | -35.915
25.5891 | -35.5739 | -35.92
25.5952 | -35.5743 | -35.9247
25.6013 | -35.5748 | -35.9292
25.6074 | -35.5753 | -35.9338
25.6135 | -35.5756 | -35.9395
25.6196 | -35.5761 | -35.9449
25.6257 | -35.5765 | -35.9506
25.6318 | -35.577 | -35.9562
25.6379 | -35.5774 | -35.9615
25.644 | -35.5778 | -35.9664
25.6501 | -35.5782 | -35.972
25.6562 | -35.5786 | -35.9774
25.6623 | -35.579 | -35.9828
25.6684 | -35.5795 | -35.9882
25.6745 | -35.5799 | -35.9935
25.6806 | -35.5803 | -35.9981
25.6867 | -35.5807 | -36.0034
25.6928 | -35.5812 | -36.0088
25.6989 | -35.5815 | -36.014
25.705 | -35.5819 | -36.0196
25.7111 | -35.5824 | -36.025
25.7173 | -35.5829 | -36.0305
25.7234 | -35.5833 | -36.0351
25.7295 | -35.5837 | -36.0403
25.7356 | -35.5841 | -36.0458
25.7417 | -35.5845 | -36.0513
25.7478 | -35.5849 | -36.0558
25.7539 | -35.5853 | -36.0609
25.76 | -35.5857 | -36.0665
25.7661 | -35.5861 | -36.0722
25.7722 | -35.5866 | -36.0776
25.7783 | -35.587 | -36.083
25.7844 | -35.5874 | -36.0884
25.7905 | -35.5878 | -36.0943
25.7966 | -35.5882 | -36.0994
25.8027 | -35.5886 | -36.1048
25.8088 | -35.589 | -36.11
25.8149 | -35.5894 | -36.1153
25.821 | -35.5898 | -36.1203
25.8271 | -35.5903 | -36.1256
25.8332 | -35.5907 | -36.1302
25.8393 | -35.5912 | -36.1354
25.8454 | -35.5916 | -36.1412
25.8515 | -35.592 | -36.1468
25.8576 | -35.5924 | -36.1518
25.8638 | -35.5928 | -36.157
25.8699 | -35.5932 | -36.1623
25.876 | -35.5936 | -36.1672
25.8821 | -35.594 | -36.1728
25.8882 | -35.5944 | -36.1779
25.8943 | -35.5948 | -36.1832
25.9004 | -35.5952 | -36.1885
25.9065 | -35.5956 | -36.1941
25.9126 | -35.596 | -36.1994
25.9187 | -35.5964 | -36.2047
25.9248 | -35.5967 | -36.2094
25.9309 | -35.5971 | -36.2149
25.937 | -35.5976 | -36.2201
25.9431 | -35.598 | -36.2251
25.9492 | -35.5984 | -36.2306
25.9553 | -35.5988 | -36.2362
25.9614 | -35.599 | -36.2416
25.9675 | -35.5995 | -36.2466
25.9736 | -35.5999 | -36.252
25.9797 | -35.6003 | -36.257
25.9858 | -35.6007 | -36.2621
25.9919 | -35.6011 | -36.2679
25.998 | -35.6015 | -36.2729
26.0042 | -35.602 | -36.2781
26.0103 | -35.6023 | -36.2835
26.0164 | -35.6028 | -36.2884
26.0225 | -35.6032 | -36.2939
26.0286 | -35.6036 | -36.2992
26.0347 | -35.6041 | -36.3041
26.0408 | -35.6045 | -36.3095
26.0469 | -35.6049 | -36.315
26.053 | -35.6053 | -36.3205
26.0591 | -35.6057 | -36.3257
26.0652 | -35.606 | -36.3304
26.0713 | -35.6064 | -36.3354
26.0774 | -35.6067 | -36.3408
26.0835 | -35.6071 | -36.3461
26.0896 | -35.6075 | -36.3511
26.0957 | -35.6078 | -36.3563
26.1018 | -35.6082 | -36.3618
26.1079 | -35.6086 | -36.3669
26.114 | -35.609 | -36.3721
26.1201 | -35.6094 | -36.3773
26.1262 | -35.6098 | -36.3822
26.1323 | -35.6101 | -36.3875
26.1384 | -35.6105 | -36.3921
26.1445 | -35.6108 | -36.3976
26.1507 | -35.6112 | -36.4028
26.1568 | -35.6116 | -36.4081
26.1629 | -35.612 | -36.4139
26.169 | -35.6124 | -36.419
26.1751 | -35.6127 | -36.4245
26.1812 | -35.6131 | -36.4299
26.1873 | -35.6135 | -36.435
26.1934 | -35.6139 | -36.4403
26.1995 | -35.6144 | -36.4454
26.2056 | -35.6147 | -36.451
26.2117 | -35.6151 | -36.4562
26.2178 | -35.6154 | -36.4617
26.2239 | -35.6158 | -36.4671
26.23 | -35.6162 | -36.4721
26.2361 | -35.6166 | -36.4769
26.2422 | -35.6169 | -36.4825
26.2483 | -35.6173 | -36.4878
26.2544 | -35.6177 | -36.4928
26.2605 | -35.618 | -36.4979
26.2666 | -35.6185 | -36.5032
26.2727 | -35.6188 | -36.5086
26.2788 | -35.6191 | -36.5139
26.2849 | -35.6194 | -36.519
26.2911 | -35.6199 | -36.5243
26.2972 | -35.6203 | -36.5294
26.3033 | -35.6207 | -36.5349
26.3094 | -35.6211 | -36.54
26.3155 | -35.6215 | -36.5451
26.3216 | -35.6219 | -36.55
26.3277 | -35.6223 | -36.5557
26.3338 | -35.6226 | -36.5613
26.3399 | -35.623 | -36.5661
26.346 | -35.6234 | -36.5708
26.3521 | -35.6238 | -36.5761
26.3582 | -35.6241 | -36.5813
26.3643 | -35.6245 | -36.5869
26.3704 | -35.6249 | -36.592
26.3765 | -35.6252 | -36.5972
26.3826 | -35.6255 | -36.6021
26.3887 | -35.6259 | -36.6075
26.3948 | -35.6263 | -36.6126
26.4009 | -35.6267 | -36.6185
26.407 | -35.6271 | -36.6244
26.4131 | -35.6274 | -36.6298
26.4192 | -35.6278 | -36.6346
26.4253 | -35.6281 | -36.6394
26.4314 | -35.6285 | -36.645
26.4376 | -35.6289 | -36.6501
26.4437 | -35.6293 | -36.6555
26.4498 | -35.6297 | -36.6607
26.4559 | -35.63 | -36.6662
26.462 | -35.6304 | -36.6716
26.4681 | -35.6307 | -36.6768
26.4742 | -35.631 | -36.682
26.4803 | -35.6314 | -36.687
26.4864 | -35.6317 | -36.6924
26.4925 | -35.6321 | -36.6977
26.4986 | -35.6324 | -36.7027
26.5047 | -35.6328 | -36.7078
26.5108 | -35.6332 | -36.713
26.5169 | -35.6336 | -36.7182
26.523 | -35.6339 | -36.7231
26.5291 | -35.6343 | -36.7283
26.5352 | -35.6347 | -36.7333
26.5413 | -35.6351 | -36.7387
26.5474 | -35.6355 | -36.7437
26.5535 | -35.6358 | -36.749
26.5596 | -35.6361 | -36.7542
26.5657 | -35.6365 | -36.7594
26.5718 | -35.6369 | -36.7647
26.578 | -35.6372 | -36.7698
26.5841 | -35.6375 | -36.7749
26.5902 | -35.6379 | -36.7796
26.5963 | -35.6382 | -36.7849
26.6024 | -35.6385 | -36.7903
26.6085 | -35.6389 | -36.7957
26.6146 | -35.6393 | -36.8012
26.6207 | -35.6397 | -36.8058
26.6268 | -35.64 | -36.811
26.6329 | -35.6403 | -36.8157
26.639 | -35.6408 | -36.8209
26.6451 | -35.6413 | -36.8258
26.6512 | -35.6416 | -36.8307
26.6573 | -35.642 | -36.8361
26.6634 | -35.6424 | -36.841
26.6695 | -35.6427 | -36.8464
26.6756 | -35.6431 | -36.8513
26.6817 | -35.6435 | -36.8566
26.6878 | -35.6438 | -36.8622
26.6939 | -35.6442 | -36.8676
26.7 | -35.6445 | -36.8721
26.7061 | -35.6448 | -36.8774
26.7122 | -35.6452 | -36.8829
26.7183 | -35.6455 | -36.888
26.7245 | -35.6458 | -36.8933
26.7306 | -35.6462 | -36.8987
26.7367 | -35.6465 | -36.9037
26.7428 | -35.6468 | -36.9092
26.7489 | -35.6472 | -36.9144
26.755 | -35.6476 | -36.9198
26.7611 | -35.6479 | -36.9252
26.7672 | -35.6482 | -36.9305
26.7733 | -35.6485 | -36.9357
26.7794 | -35.6489 | -36.9407
26.7855 | -35.6492 | -36.9461
26.7916 | -35.6496 | -36.9515
26.7977 | -35.65 | -36.9571
26.8038 | -35.6503 | -36.9614
26.8099 | -35.6506 | -36.9666
26.816 | -35.6509 | -36.9719
26.8221 | -35.6513 | -36.9771
26.8282 | -35.6517 | -36.982
26.8343 | -35.6521 | -36.9874
26.8404 | -35.6524 | -36.9923
26.8465 | -35.6528 | -36.9977
26.8526 | -35.6532 | -37.0035
26.8587 | -35.6536 | -37.0089
26.8649 | -35.6539 | -37.0136
26.871 | -35.6543 | -37.0183
26.8771 | -35.6546 | -37.0235
26.8832 | -35.655 | -37.0286
26.8893 | -35.6554 | -37.0344
26.8954 | -35.6557 | -37.0393
26.9015 | -35.6561 | -37.0444
26.9076 | -35.6565 | -37.0495
26.9137 | -35.6567 | -37.0548
26.9198 | -35.657 | -37.0598
26.9259 | -35.6574 | -37.0654
26.932 | -35.6577 | -37.0709
26.9381 | -35.658 | -37.0757
26.9442 | -35.6583 | -37.0813
26.9503 | -35.6587 | -37.0861
26.9564 | -35.659 | -37.0908
26.9625 | -35.6593 | -37.0963
26.9686 | -35.6597 | -37.101
26.9747 | -35.66 | -37.1064
26.9808 | -35.6603 | -37.1114
26.9869 | -35.6607 | -37.1167
26.993 | -35.661 | -37.1218
26.9991 | -35.6613 | -37.1271
27.0052 | -35.6617 | -37.1324
27.0114 | -35.662 | -37.1382
27.0175 | -35.6624 | -37.1436
27.0236 | -35.6627 | -37.149
27.0297 | -35.663 | -37.1539
27.0358 | -35.6634 | -37.1586
27.0419 | -35.6637 | -37.1636
27.048 | -35.6641 | -37.1686
27.0541 | -35.6645 | -37.1737
27.0602 | -35.6648 | -37.1788
27.0663 | -35.6651 | -37.1834
27.0724 | -35.6655 | -37.1889
27.0785 | -35.6658 | -37.1939
27.0846 | -35.6662 | -37.1983
27.0907 | -35.6665 | -37.2034
27.0968 | -35.6669 | -37.2083
27.1029 | -35.6672 | -37.2134
27.109 | -35.6676 | -37.2187
27.1151 | -35.6678 | -37.2237
27.1212 | -35.6682 | -37.2286
27.1273 | -35.6685 | -37.2337
27.1334 | -35.6688 | -37.2387
27.1395 | -35.6691 | -37.2444
27.1456 | -35.6694 | -37.2497
27.1518 | -35.6698 | -37.2548
27.1579 | -35.67 | -37.259
27.164 | -35.6704 | -37.2633
27.1701 | -35.6707 | -37.2684
27.1762 | -35.6711 | -37.2735
27.1823 | -35.6714 | -37.2789
27.1884 | -35.6718 | -37.2843
27.1945 | -35.6721 | -37.2894
27.2006 | -35.6724 | -37.2945
27.2067 | -35.6728 | -37.2997
27.2128 | -35.673 | -37.3049
27.2189 | -35.6734 | -37.3105
27.225 | -35.6736 | -37.3159
27.2311 | -35.674 | -37.321
27.2372 | -35.6743 | -37.3258
27.2433 | -35.6746 | -37.3313
27.2494 | -35.675 | -37.3365
27.2555 | -35.6754 | -37.3415
27.2616 | -35.6757 | -37.3461
27.2677 | -35.676 | -37.3516
27.2738 | -35.6764 | -37.356
27.2799 | -35.6767 | -37.3604
27.286 | -35.6769 | -37.3653
27.2921 | -35.6773 | -37.3705
27.2983 | -35.6777 | -37.3762
27.3044 | -35.678 | -37.3808
27.3105 | -35.6783 | -37.3857
27.3166 | -35.6787 | -37.3915
27.3227 | -35.679 | -37.3965
27.3288 | -35.6793 | -37.4015
27.3349 | -35.6796 | -37.4066
27.341 | -35.6799 | -37.411
27.3471 | -35.6801 | -37.4163
27.3532 | -35.6805 | -37.4213
27.3593 | -35.6807 | -37.4265
27.3654 | -35.681 | -37.4323
27.3715 | -35.6813 | -37.4371
27.3776 | -35.6817 | -37.4423
27.3837 | -35.682 | -37.4472
27.3898 | -35.6823 | -37.4522
27.3959 | -35.6826 | -37.4575
27.402 | -35.6828 | -37.4623
27.4081 | -35.6832 | -37.4675
27.4142 | -35.6835 | -37.4724
27.4203 | -35.6838 | -37.4775
27.4264 | -35.6841 | -37.4831
27.4325 | -35.6845 | -37.4878
27.4387 | -35.6848 | -37.4928
27.4448 | -35.6851 | -37.4986
27.4509 | -35.6853 | -37.5039
27.457 | -35.6857 | -37.5092
27.4631 | -35.6861 | -37.5138
27.4692 | -35.6864 | -37.5192
27.4753 | -35.6866 | -37.5235
27.4814 | -35.687 | -37.529
27.4875 | -35.6873 | -37.5343
27.4936 | -35.6876 | -37.5394
27.4997 | -35.6879 | -37.5448
27.5058 | -35.6882 | -37.5504
27.5119 | -35.6886 | -37.5559
27.518 | -35.6889 | -37.5603
27.5241 | -35.6893 | -37.5651
27.5302 | -35.6896 | -37.5702
27.5363 | -35.6899 | -37.5754
27.5424 | -35.6902 | -37.5808
27.5485 | -35.6906 | -37.5861
27.5546 | -35.6909 | -37.5909
27.5607 | -35.6911 | -37.5958
27.5668 | -35.6915 | -37.6011
27.5729 | -35.6919 | -37.6063
27.5791 | -35.6922 | -37.6111
27.5852 | -35.6924 | -37.6167
27.5913 | -35.6927 | -37.622
27.5974 | -35.693 | -37.627
27.6035 | -35.6933 | -37.6319
27.6096 | -35.6936 | -37.6371
27.6157 | -35.6939 | -37.6424
27.6218 | -35.6942 | -37.6478
27.6279 | -35.6945 | -37.6531
27.634 | -35.6949 | -37.6584
27.6401 | -35.6953 | -37.6637
27.6462 | -35.6956 | -37.6686
27.6523 | -35.6959 | -37.6732
27.6584 | -35.6962 | -37.6784
27.6645 | -35.6965 | -37.6837
27.6706 | -35.6968 | -37.688
27.6767 | -35.6971 | -37.6927
27.6828 | -35.6974 | -37.6979
27.6889 | -35.6977 | -37.7032
27.695 | -35.698 | -37.7081
27.7011 | -35.6983 | -37.7128
27.7072 | -35.6986 | -37.7176
27.7133 | -35.6988 | -37.7227
27.7194 | -35.699 | -37.7274
27.7256 | -35.6994 | -37.7327
27.7317 | -35.6996 | -37.7374
27.7378 | -35.6999 | -37.7418
27.7439 | -35.7002 | -37.7468
27.75 | -35.7004 | -37.7519
27.7561 | -35.7008 | -37.7568
27.7622 | -35.7011 | -37.7625
27.7683 | -35.7015 | -37.7669
27.7744 | -35.7019 | -37.7722
27.7805 | -35.7022 | -37.7777
27.7866 | -35.7025 | -37.7827
27.7927 | -35.7029 | -37.7876
27.7988 | -35.7032 | -37.7918
27.8049 | -35.7034 | -37.7969
27.811 | -35.7038 | -37.8026
27.8171 | -35.7041 | -37.8076
27.8232 | -35.7044 | -37.8127
27.8293 | -35.7048 | -37.8176
27.8354 | -35.7052 | -37.8227
27.8415 | -35.7054 | -37.8276
27.8476 | -35.7058 | -37.8325
27.8537 | -35.7062 | -37.8371
27.8598 | -35.7065 | -37.8411
27.866 | -35.7068 | -37.8465
27.8721 | -35.707 | -37.8518
27.8782 | -35.7073 | -37.8566
27.8843 | -35.7076 | -37.8612
27.8904 | -35.7079 | -37.8663
27.8965 | -35.7082 | -37.8711
27.9026 | -35.7084 | -37.8756
27.9087 | -35.7088 | -37.8806
27.9148 | -35.7091 | -37.8851
27.9209 | -35.7094 | -37.89
27.927 | -35.7097 | -37.8952
27.9331 | -35.71 | -37.8993
27.9392 | -35.7102 | -37.9037
27.9453 | -35.7105 | -37.9089
27.9514 | -35.7109 | -37.9138
27.9575 | -35.7111 | -37.9194
27.9636 | -35.7113 | -37.9244
27.9697 | -35.7116 | -37.9289
27.9758 | -35.7119 | -37.9339
27.9819 | -35.7123 | -37.9387
27.988 | -35.7125 | -37.9441
27.9941 | -35.7128 | -37.949
28.0002 | -35.7131 | -37.9536
28.0063 | -35.7134 | -37.9587
28.0125 | -35.7136 | -37.9641
28.0186 | -35.7139 | -37.969
28.0247 | -35.7141 | -37.9738
28.0308 | -35.7144 | -37.9787
28.0369 | -35.7146 | -37.9837
28.043 | -35.7149 | -37.9885
28.0491 | -35.7151 | -37.9929
28.0552 | -35.7155 | -37.9979
28.0613 | -35.7158 | -38.0029
28.0674 | -35.7161 | -38.0078
28.0735 | -35.7163 | -38.0138

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5253 | -41.8083
21.4908 | -35.5581 | -35.0487
16.4948 | -34.9226 | -27.6898
12.6602 | -34.5519 | -21.4214
9.7171 | -34.3235 | -16.5169
7.4582 | -34.1687 | -12.5285
5.7244 | -34.0618 | -9.2958
4.3936 | -33.9919 | -6.6891
3.3722 | -33.977 | -4.6364
2.5883 | -33.9732 | -3.0772
1.9866 | -33.9648 | -1.8358
1.5248 | -33.963 | -0.83391
1.1703 | -33.9684 | -0.18507
1 | -33.9664 | 0.13998
0.89822 | -33.9716 | 0.40926
0.68941 | -33.9381 | 0.99206
0.52919 | -33.9433 | 1.4403
0.40612 | -33.9381 | 2.0482
0.31171 | -33.9452 | 2.4961
0.23928 | -33.9458 | 3.2268
0.18358 | -33.9475 | 3.6517
0.14091 | -33.9621 | 4.5285
0.10819 | -33.9957 | 5.6867
0.083017 | -33.9875 | 7.1704
0.063705 | -34.0515 | 9.0344
0.048923 | -34.1121 | 11.5294
0.037479 | -34.2207 | 14.7142
0.028801 | -34.3966 | 18.7225
0.022078 | -34.6955 | 24.0903
0.017023 | -35.1769 | 30.5216
0.013018 | -35.9768 | 38.1897
0.010014 | -37.0968 | 46.1098

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003719 | -1.3487
0.007438 | -24.8
0.011158 | -32.8961
0.014877 | -37.7552
0.018596 | -40.1412
0.022315 | -42.3002
0.026034 | -45.913
0.029754 | -46.7023
0.033473 | -47.3413
0.037192 | -49.0871
0.040911 | -50.5925
0.04463 | -51.7961
0.04835 | -52.9434
0.052069 | -54.1458
0.055788 | -54.2648
0.059507 | -55.2383
0.063226 | -55.8272
0.066946 | -57.5182
0.070665 | -59.3275
0.074384 | -59.7069
0.078103 | -60.1652
0.081822 | -60.1418
0.085542 | -60.4259
0.089261 | -60.5864
0.09298 | -60.806
0.096699 | -62.2297
0.10042 | -63.0921
0.10414 | -63.2286
0.10786 | -62.6739
0.11158 | -63.6615
0.11529 | -64.6075
0.11901 | -64.4444
0.12273 | -64.4456
0.12645 | -65.1589
0.13017 | -65.522
0.13389 | -65.8023
0.13761 | -66.3556
0.14133 | -67.1437
0.14505 | -67.3812
0.14877 | -67.6751
0.15249 | -67.2439
0.15621 | -68.1358
0.15993 | -67.7474
0.16365 | -66.9048
0.16736 | -67.2743
0.17108 | -67.7864
0.1748 | -68.7754
0.17852 | -69.4316
0.18224 | -69.3633
0.18596 | -69.8479
0.18968 | -69.5779
0.1934 | -70.0466
0.19712 | -71.4823
0.20084 | -70.4632
0.20456 | -70.9383
0.20827 | -70.5384
0.21199 | -71.0169
0.21571 | -71.824
0.21943 | -71.8254
0.22315 | -72.7833
0.22687 | -73.0459
0.23059 | -72.1696
0.23431 | -71.3485
0.23803 | -70.638
0.24175 | -70.8389
0.24547 | -72.8422
0.24919 | -73.2947
0.25291 | -72.7423
0.25662 | -73.3858
0.26034 | -74.1012
0.26406 | -73.5043
0.26778 | -73.5149
0.2715 | -73.3309
0.27522 | -74.0969
0.27894 | -73.6237
0.28266 | -74.1957
0.28638 | -74.8565
0.2901 | -74.2015
0.29382 | -73.3658
0.29754 | -74.3003
0.30125 | -74.5076
0.30497 | -74.5869
0.30869 | -75.6242
0.31241 | -75.0462
0.31613 | -74.4762
0.31985 | -74.2439
0.32357 | -75.3777
0.32729 | -75.6547
0.33101 | -75.5076
0.33473 | -75.8548
0.33845 | -76.0771
0.34217 | -75.3121
0.34589 | -75.8972
0.3496 | -77.251
0.35332 | -76.7878
0.35704 | -76.7894
0.36076 | -77.2305
0.36448 | -77.1248
0.3682 | -76.4627
0.37192 | -76.4204
0.37564 | -76.7678
0.37936 | -76.0954
0.38308 | -76.849
0.3868 | -78.2095
0.39052 | -78.5109
0.39424 | -77.7202
0.39795 | -77.4235
0.40167 | -77.0404
0.40539 | -76.6333
0.40911 | -76.6167
0.41283 | -77.9577
0.41655 | -78.479
0.42027 | -78.5526
0.42399 | -77.8492
0.42771 | -78.056
0.43143 | -78.8569
0.43515 | -78.6979
0.43887 | -78.7522
0.44259 | -78.2874
0.4463 | -77.6598
0.45002 | -78.5287
0.45374 | -79.5941
0.45746 | -78.755
0.46118 | -79.5321
0.4649 | -80.184
0.46862 | -80.0248
0.47234 | -80.0131
0.47606 | -80.5349
0.47978 | -80.8048
0.4835 | -79.6391
0.48722 | -79.9065
0.49093 | -79.3447
0.49465 | -78.9622
0.49837 | -79.5487
0.50209 | -79.8242
0.50581 | -80.3564
0.50953 | -79.8443
0.51325 | -79.8048
0.51697 | -79.3582
0.52069 | -78.8279
0.52441 | -79.4035
0.52813 | -79.4305
0.53185 | -80.7993
0.53556 | -81.1122
0.53928 | -80.8465
0.543 | -80.3787
0.54672 | -80.2536
0.55044 | -80.2458
0.55416 | -80.2033
0.55788 | -80.4468
0.5616 | -80.8407
0.56532 | -81.0118
0.56904 | -81.065
0.57276 | -81.2255
0.57648 | -81.6218
0.5802 | -81.7872
0.58391 | -81.8749
0.58763 | -81.8805
0.59135 | -81.6955
0.59507 | -81.2624
0.59879 | -80.8748
0.60251 | -81.1391
0.60623 | -81.6253
0.60995 | -81.5846
0.61367 | -81.6167
0.61739 | -81.7407
0.62111 | -81.6343
0.62483 | -81.7772
0.62855 | -82.1072
0.63226 | -82.2735
0.63598 | -82.3316
0.6397 | -82.5133
0.64342 | -82.7205
0.64714 | -82.345
0.65086 | -81.677
0.65458 | -81.653
0.6583 | -82.0759
0.66202 | -82.4619
0.66574 | -82.7013
0.66946 | -82.7299
0.67317 | -82.6456
0.67689 | -82.652
0.68061 | -82.51
0.68433 | -82.5792
0.68805 | -83.0271
0.69177 | -83.7212
0.69549 | -84.2997
0.69921 | -83.9233
0.70293 | -83.0285
0.70665 | -82.8847
0.71037 | -83.4086
0.71409 | -84.0098
0.71781 | -83.7918
0.72152 | -83.6643
0.72524 | -84.2413
0.72896 | -84.557
0.73268 | -83.9335
0.7364 | -82.9789
0.74012 | -82.8058
0.74384 | -83.691
0.74756 | -84.649
0.75128 | -84.5287
0.755 | -83.8088
0.75872 | -83.6237
0.76244 | -83.8947
0.76616 | -84.1754
0.76987 | -84.1439
0.77359 | -83.8491
0.77731 | -83.6047
0.78103 | -83.8687
0.78475 | -84.2657
0.78847 | -84.4003
0.79219 | -84.1636
0.79591 | -84.0835
0.79963 | -84.5
0.80335 | -84.8659
0.80707 | -84.8524
0.81079 | -84.7758
0.81451 | -84.4592
0.81822 | -84.09
0.82194 | -83.9796
0.82566 | -84.0972
0.82938 | -84.4021
0.8331 | -85.2032
0.83682 | -85.7227
0.84054 | -85.3753
0.84426 | -84.8314
0.84798 | -84.3584
0.8517 | -84.0368
0.85542 | -84.3139
0.85913 | -84.7448
0.86285 | -84.736
0.86657 | -85.1084
0.87029 | -85.5368
0.87401 | -85.3833
0.87773 | -85.1003
0.88145 | -85.1718
0.88517 | -85.2772
0.88889 | -85.0407
0.89261 | -85.0574
0.89633 | -85.2531
0.90005 | -85.2474
0.90377 | -85.2748
0.90748 | -85.4962
0.9112 | -85.7292
0.91492 | -85.6655
0.91864 | -85.7529
0.92236 | -85.8624
0.92608 | -85.5977
0.9298 | -85.6281
0.93352 | -86.0544
0.93724 | -86.0367
0.94096 | -85.8498
0.94468 | -86.3173
0.9484 | -86.8447
0.95212 | -86.6745
0.95583 | -86.1088
0.95955 | -85.6733
0.96327 | -85.8315
0.96699 | -86.1774
0.97071 | -86.3436
0.97443 | -86.3337
0.97815 | -86.3623
0.98187 | -86.2881
0.98559 | -86.0812
0.98931 | -85.5539
0.99303 | -85.3287
0.99675 | -85.1861
1.0005 | -85.2587
1.0042 | -86.3494
1.0079 | -87.0063
1.0116 | -86.8436
1.0153 | -87.0309
1.0191 | -87.4321
1.0228 | -87.6787
1.0265 | -87.5952
1.0302 | -87.2406
1.0339 | -86.9626
1.0377 | -86.8679
1.0414 | -86.976
1.0451 | -87.1224
1.0488 | -87.0918
1.0525 | -86.7456
1.0563 | -86.4906
1.06 | -86.6227
1.0637 | -86.8041
1.0674 | -86.8844
1.0711 | -87.017
1.0748 | -87.0647
1.0786 | -87.0199
1.0823 | -86.9144
1.086 | -86.8352
1.0897 | -86.953
1.0934 | -87.2983
1.0972 | -87.6514
1.1009 | -87.853
1.1046 | -87.9482
1.1083 | -88.0563
1.112 | -88.1999
1.1158 | -88.21
1.1195 | -87.9288
1.1232 | -87.5328
1.1269 | -87.3866
1.1306 | -87.6077
1.1344 | -87.8928
1.1381 | -87.9621
1.1418 | -87.9496
1.1455 | -88.0351
1.1492 | -88.2395
1.153 | -88.4895
1.1567 | -88.5545
1.1604 | -88.4025
1.1641 | -88.327
1.1678 | -88.4096
1.1715 | -88.4711
1.1753 | -88.4052
1.179 | -88.2696
1.1827 | -88.2165
1.1864 | -88.288
1.1901 | -88.4396
1.1939 | -88.6467
1.1976 | -88.7187
1.2013 | -88.4701
1.205 | -88.1941
1.2087 | -88.2582
1.2125 | -88.636
1.2162 | -88.9905
1.2199 | -89.0366
1.2236 | -88.8557
1.2273 | -88.6403
1.2311 | -88.4521
1.2348 | -88.3323
1.2385 | -88.3381
1.2422 | -88.5653
1.2459 | -88.968
1.2497 | -89.2912
1.2534 | -89.411
1.2571 | -89.4172
1.2608 | -89.4609
1.2645 | -89.5963
1.2682 | -89.6061
1.272 | -89.3674
1.2757 | -89.1284
1.2794 | -89.0693
1.2831 | -89.1186
1.2868 | -89.2227
1.2906 | -89.3892
1.2943 | -89.591
1.298 | -89.7234
1.3017 | -89.7269
1.3054 | -89.6766
1.3092 | -89.5886
1.3129 | -89.5401
1.3166 | -89.6325
1.3203 | -89.806
1.324 | -89.9055
1.3278 | -89.8186
1.3315 | -89.6217
1.3352 | -89.4633
1.3389 | -89.3853
1.3426 | -89.4675
1.3464 | -89.7183
1.3501 | -89.887
1.3538 | -89.7729
1.3575 | -89.6169
1.3612 | -89.6197
1.3649 | -89.6292
1.3687 | -89.4717
1.3724 | -89.1909
1.3761 | -89.0244
1.3798 | -89.223
1.3835 | -89.7722
1.3873 | -90.3399
1.391 | -90.463
1.3947 | -90.1302
1.3984 | -89.8537
1.4021 | -89.9297
1.4059 | -90.2131
1.4096 | -90.3189
1.4133 | -90.1172
1.417 | -89.9355
1.4207 | -89.9703
1.4245 | -90.1014
1.4282 | -90.3378
1.4319 | -90.7046
1.4356 | -90.8557
1.4393 | -90.5852
1.443 | -90.2021
1.4468 | -89.9409
1.4505 | -89.7852
1.4542 | -89.7234
1.4579 | -89.8738
1.4616 | -90.2163
1.4654 | -90.5126
1.4691 | -90.5987
1.4728 | -90.6712
1.4765 | -90.9268
1.4802 | -91.1212
1.484 | -90.8877
1.4877 | -90.3174
1.4914 | -89.7992
1.4951 | -89.5947
1.4988 | -89.6989
1.5026 | -89.9014
1.5063 | -89.9661
1.51 | -89.9463
1.5137 | -90.0683
1.5174 | -90.3398
1.5212 | -90.637
1.5249 | -90.8572
1.5286 | -90.869
1.5323 | -90.784
1.536 | -90.8484
1.5397 | -90.9955
1.5435 | -90.9636
1.5472 | -90.7118
1.5509 | -90.4962
1.5546 | -90.5924
1.5583 | -90.9572
1.5621 | -91.2719
1.5658 | -91.4296
1.5695 | -91.6175
1.5732 | -91.731
1.5769 | -91.5499
1.5807 | -91.2783
1.5844 | -91.1969
1.5881 | -91.2472
1.5918 | -91.1837
1.5955 | -91.0232
1.5993 | -91.0002
1.603 | -91.161
1.6067 | -91.4251
1.6104 | -91.6214
1.6141 | -91.7133
1.6179 | -91.6923
1.6216 | -91.6631
1.6253 | -91.6785
1.629 | -91.7027
1.6327 | -91.6827
1.6364 | -91.6187
1.6402 | -91.5739
1.6439 | -91.5635
1.6476 | -91.4884
1.6513 | -91.2114
1.655 | -90.794
1.6588 | -90.4585
1.6625 | -90.3447
1.6662 | -90.4401
1.6699 | -90.6809
1.6736 | -91.0069
1.6774 | -91.334
1.6811 | -91.4917
1.6848 | -91.3706
1.6885 | -91.0502
1.6922 | -90.7099
1.696 | -90.508
1.6997 | -90.5031
1.7034 | -90.655
1.7071 | -90.9034
1.7108 | -91.2199
1.7146 | -91.5576
1.7183 | -91.8075
1.722 | -91.9108
1.7257 | -91.9117
1.7294 | -91.8939
1.7331 | -91.8804
1.7369 | -91.8421
1.7406 | -91.7778
1.7443 | -91.7341
1.748 | -91.721
1.7517 | -91.7172
1.7555 | -91.7248
1.7592 | -91.7393
1.7629 | -91.7405
1.7666 | -91.7589
1.7703 | -91.8699
1.7741 | -92.0864
1.7778 | -92.3057
1.7815 | -92.4369
1.7852 | -92.4641
1.7889 | -92.4577
1.7927 | -92.4409
1.7964 | -92.3161
1.8001 | -92.048
1.8038 | -91.7221
1.8075 | -91.4917
1.8113 | -91.486
1.815 | -91.7451
1.8187 | -92.1241
1.8224 | -92.3531
1.8261 | -92.3076
1.8298 | -92.1182
1.8336 | -91.9537
1.8373 | -91.9031
1.841 | -91.9835
1.8447 | -92.1347
1.8484 | -92.2132
1.8522 | -92.1445
1.8559 | -91.9902
1.8596 | -91.8827
1.8633 | -91.9097
1.867 | -92.0927
1.8708 | -92.4007
1.8745 | -92.7553
1.8782 | -93.0977
1.8819 | -93.4086
1.8856 | -93.6535
1.8894 | -93.7817
1.8931 | -93.7981
1.8968 | -93.7188
1.9005 | -93.5851
1.9042 | -93.4772
1.908 | -93.482
1.9117 | -93.5866
1.9154 | -93.666
1.9191 | -93.6182
1.9228 | -93.4489
1.9265 | -93.2277
1.9303 | -93.0479
1.934 | -92.9549
1.9377 | -92.9513
1.9414 | -93.0311
1.9451 | -93.1935
1.9489 | -93.3845
1.9526 | -93.4216
1.9563 | -93.2257
1.96 | -92.9469
1.9637 | -92.812
1.9675 | -92.9212
1.9712 | -93.2229
1.9749 | -93.5238
1.9786 | -93.6212
1.9823 | -93.4746
1.9861 | -93.219
1.9898 | -92.9199
1.9935 | -92.649
1.9972 | -92.4852
2.0009 | -92.4691
2.0046 | -92.5608
2.0084 | -92.7275
2.0121 | -92.9388
2.0158 | -93.152
2.0195 | -93.2925
2.0232 | -93.3146
2.027 | -93.2173
2.0307 | -93.0661
2.0344 | -92.9195
2.0381 | -92.7999
2.0418 | -92.7401
2.0456 | -92.7502
2.0493 | -92.8035
2.053 | -92.8501
2.0567 | -92.9057
2.0604 | -93.0463
2.0642 | -93.3167
2.0679 | -93.6224
2.0716 | -93.8048
2.0753 | -93.7725
2.079 | -93.5759
2.0828 | -93.3279
2.0865 | -93.1982
2.0902 | -93.2804
2.0939 | -93.4945
2.0976 | -93.6466
2.1013 | -93.6086
2.1051 | -93.42
2.1088 | -93.2037
2.1125 | -93.0424
2.1162 | -92.9832
2.1199 | -93.0574
2.1237 | -93.2543
2.1274 | -93.5053
2.1311 | -93.7429
2.1348 | -93.961
2.1385 | -94.1782
2.1423 | -94.3695
2.146 | -94.4377
2.1497 | -94.4603
2.1534 | -94.3865
2.1571 | -94.2319
2.1609 | -94.0449
2.1646 | -93.8937
2.1683 | -93.81
2.172 | -93.7682
2.1757 | -93.7205
2.1795 | -93.6392
2.1832 | -93.5586
2.1869 | -93.5218
2.1906 | -93.5156
2.1943 | -93.5027
2.198 | -93.4304
2.2018 | -93.2979
2.2055 | -93.1679
2.2092 | -93.1108
2.2129 | -93.1571
2.2166 | -93.2809
2.2204 | -93.4447
2.2241 | -93.5941
2.2278 | -93.7262
2.2315 | -93.8456
2.2352 | -93.9483
2.239 | -94.0322
2.2427 | -94.082
2.2464 | -94.057
2.2501 | -93.9467
2.2538 | -93.7793
2.2576 | -93.6216
2.2613 | -93.5469
2.265 | -93.5977
2.2687 | -93.746
2.2724 | -93.9276
2.2762 | -94.0534
2.2799 | -94.0925
2.2836 | -94.0608
2.2873 | -94.0087
2.291 | -93.9907
2.2947 | -94.0377
2.2985 | -94.138
2.3022 | -94.2452
2.3059 | -94.316
2.3096 | -94.3381
2.3133 | -94.3415
2.3171 | -94.3563
2.3208 | -94.4035
2.3245 | -94.4571
2.3282 | -94.481
2.3319 | -94.4259
2.3357 | -94.2784
2.3394 | -94.0499
2.3431 | -93.7865
2.3468 | -93.546
2.3505 | -93.3853
2.3543 | -93.3387
2.358 | -93.4193
2.3617 | -93.6067
2.3654 | -93.8796
2.3691 | -94.1636
2.3729 | -94.3763
2.3766 | -94.4782
2.3803 | -94.478
2.384 | -94.3988
2.3877 | -94.3027
2.3914 | -94.2585
2.3952 | -94.2839
2.3989 | -94.3357
2.4026 | -94.3539
2.4063 | -94.3283
2.41 | -94.309
2.4138 | -94.3368
2.4175 | -94.4071
2.4212 | -94.4808
2.4249 | -94.5181
2.4286 | -94.4916
2.4324 | -94.4316
2.4361 | -94.4222
2.4398 | -94.5038
2.4435 | -94.6467
2.4472 | -94.8304
2.451 | -95.0532
2.4547 | -95.3422
2.4584 | -95.6586
2.4621 | -95.8989
2.4658 | -95.964
2.4695 | -95.8513
2.4733 | -95.6336
2.477 | -95.4159
2.4807 | -95.2618
2.4844 | -95.1915
2.4881 | -95.1351
2.4919 | -95.0496
2.4956 | -94.9264
2.4993 | -94.8206
2.503 | -94.7752
2.5067 | -94.8091
2.5105 | -94.9225
2.5142 | -95.0729
2.5179 | -95.1809
2.5216 | -95.1811
2.5253 | -95.0539
2.5291 | -94.8784
2.5328 | -94.7455
2.5365 | -94.7123
2.5402 | -94.7878
2.5439 | -94.9323
2.5477 | -95.0567
2.5514 | -95.0801
2.5551 | -94.9941
2.5588 | -94.8666
2.5625 | -94.8017
2.5662 | -94.8455
2.57 | -94.9687
2.5737 | -95.1041
2.5774 | -95.192
2.5811 | -95.2294
2.5848 | -95.2569
2.5886 | -95.3133
2.5923 | -95.409
2.596 | -95.491
2.5997 | -95.476
2.6034 | -95.3395
2.6072 | -95.1314
2.6109 | -94.9397
2.6146 | -94.8326
2.6183 | -94.8185
2.622 | -94.8693
2.6258 | -94.9682
2.6295 | -95.1084
2.6332 | -95.2727
2.6369 | -95.4215
2.6406 | -95.5024
2.6444 | -95.495
2.6481 | -95.4532
2.6518 | -95.4198
2.6555 | -95.4062
2.6592 | -95.4121
2.6629 | -95.4034
2.6667 | -95.3513
2.6704 | -95.2799
2.6741 | -95.2374
2.6778 | -95.2568
2.6815 | -95.3113
2.6853 | -95.3238
2.689 | -95.3196
2.6927 | -95.2961
2.6964 | -95.2865
2.7001 | -95.3164
2.7039 | -95.3997
2.7076 | -95.5404
2.7113 | -95.7268
2.715 | -95.921
2.7187 | -96.0759
2.7225 | -96.1579
2.7262 | -96.1486
2.7299 | -96.075
2.7336 | -95.9529
2.7373 | -95.8196
2.7411 | -95.7242
2.7448 | -95.6796
2.7485 | -95.6992
2.7522 | -95.7833
2.7559 | -95.9148
2.7596 | -96.0573
2.7634 | -96.1781
2.7671 | -96.2544
2.7708 | -96.2586
2.7745 | -96.2184
2.7782 | -96.159
2.782 | -96.0994
2.7857 | -96.05
2.7894 | -96.0102
2.7931 | -95.9978
2.7968 | -96.0104
2.8006 | -96.0503
2.8043 | -96.1108
2.808 | -96.1752
2.8117 | -96.2104
2.8154 | -96.2132
2.8192 | -96.2012
2.8229 | -96.1618
2.8266 | -96.1053
2.8303 | -96.0532
2.834 | -96.0264
2.8378 | -96.0362
2.8415 | -96.0753
2.8452 | -96.125
2.8489 | -96.1667
2.8526 | -96.1958
2.8563 | -96.2283
2.8601 | -96.2536
2.8638 | -96.2544
2.8675 | -96.2132
2.8712 | -96.1491
2.8749 | -96.0873
2.8787 | -96.0369
2.8824 | -95.9981
2.8861 | -95.9701
2.8898 | -95.9541
2.8935 | -95.9513
2.8973 | -95.9588
2.901 | -95.9623
2.9047 | -95.9711
2.9084 | -96.0206
2.9121 | -96.0936
2.9159 | -96.1779
2.9196 | -96.246
2.9233 | -96.2909
2.927 | -96.3235
2.9307 | -96.379
2.9344 | -96.4557
2.9382 | -96.5175
2.9419 | -96.5612
2.9456 | -96.581
2.9493 | -96.5897
2.953 | -96.5845
2.9568 | -96.5597
2.9605 | -96.5082
2.9642 | -96.436
2.9679 | -96.3656
2.9716 | -96.3222
2.9754 | -96.284
2.9791 | -96.211
2.9828 | -96.1191
2.9865 | -95.9784
2.9902 | -95.7947
2.994 | -95.618
2.9977 | -95.5182
3.0014 | -95.5347
3.0051 | -95.6795
3.0088 | -95.9332
3.0126 | -96.2077
3.0163 | -96.4359
3.02 | -96.561
3.0237 | -96.5942
3.0274 | -96.5687
3.0311 | -96.5097
3.0349 | -96.5055
3.0386 | -96.5628
3.0423 | -96.6651
3.046 | -96.7544
3.0497 | -96.7852
3.0535 | -96.7479
3.0572 | -96.6438
3.0609 | -96.5233
3.0646 | -96.4248
3.0683 | -96.3631
3.0721 | -96.3395
3.0758 | -96.3405
3.0795 | -96.3636
3.0832 | -96.392
3.0869 | -96.4223
3.0907 | -96.4499
3.0944 | -96.4752
3.0981 | -96.5015
3.1018 | -96.5073
3.1055 | -96.4961
3.1093 | -96.4494
3.113 | -96.3387
3.1167 | -96.1591
3.1204 | -95.937
3.1241 | -95.7433
3.1278 | -95.6476
3.1316 | -95.6778
3.1353 | -95.8196
3.139 | -96.0324
3.1427 | -96.2468
3.1464 | -96.4175
3.1502 | -96.5309
3.1539 | -96.5699
3.1576 | -96.5538
3.1613 | -96.5231
3.165 | -96.4824
3.1688 | -96.427
3.1725 | -96.3614
3.1762 | -96.3071
3.1799 | -96.2886
3.1836 | -96.3165
3.1874 | -96.3909
3.1911 | -96.4956
3.1948 | -96.598
3.1985 | -96.6884
3.2022 | -96.7633
3.206 | -96.8079
3.2097 | -96.8141
3.2134 | -96.7935
3.2171 | -96.7382
3.2208 | -96.697
3.2245 | -96.6546
3.2283 | -96.6178
3.232 | -96.5882
3.2357 | -96.5509
3.2394 | -96.5075
3.2431 | -96.4688
3.2469 | -96.4582
3.2506 | -96.4902
3.2543 | -96.5819
3.258 | -96.7245
3.2617 | -96.8989
3.2655 | -97.0772
3.2692 | -97.2374
3.2729 | -97.3538
3.2766 | -97.4027
3.2803 | -97.3991
3.2841 | -97.3602
3.2878 | -97.3178
3.2915 | -97.2812
3.2952 | -97.251
3.2989 | -97.2391
3.3026 | -97.2333
3.3064 | -97.248
3.3101 | -97.2743
3.3138 | -97.3088
3.3175 | -97.3406
3.3212 | -97.3578
3.325 | -97.3413
3.3287 | -97.3081
3.3324 | -97.2826
3.3361 | -97.2893
3.3398 | -97.3467
3.3436 | -97.4582
3.3473 | -97.6125
3.351 | -97.7806
3.3547 | -97.924
3.3584 | -98.0257
3.3622 | -98.0832
3.3659 | -98.1176
3.3696 | -98.151
3.3733 | -98.1919
3.377 | -98.2388
3.3808 | -98.2939
3.3845 | -98.318
3.3882 | -98.2994
3.3919 | -98.2253
3.3956 | -98.1065
3.3993 | -97.9475
3.4031 | -97.7827
3.4068 | -97.6476
3.4105 | -97.5634
3.4142 | -97.5339
3.4179 | -97.5479
3.4217 | -97.5811
3.4254 | -97.6043
3.4291 | -97.5998
3.4328 | -97.5687
3.4365 | -97.5112
3.4403 | -97.4542
3.444 | -97.403
3.4477 | -97.3748
3.4514 | -97.3673
3.4551 | -97.3842
3.4589 | -97.4222
3.4626 | -97.4635
3.4663 | -97.5134
3.47 | -97.5709
3.4737 | -97.6109
3.4775 | -97.6323
3.4812 | -97.6327
3.4849 | -97.6188
3.4886 | -97.5967
3.4923 | -97.5813
3.496 | -97.5704
3.4998 | -97.5736
3.5035 | -97.5962
3.5072 | -97.6326
3.5109 | -97.6791
3.5146 | -97.7291
3.5184 | -97.7771
3.5221 | -97.8011
3.5258 | -97.7872
3.5295 | -97.7566
3.5332 | -97.6987
3.537 | -97.6242
3.5407 | -97.5419
3.5444 | -97.4646
3.5481 | -97.4148
3.5518 | -97.4003
3.5556 | -97.4258
3.5593 | -97.4986
3.563 | -97.6195
3.5667 | -97.7905
3.5704 | -97.9918
3.5742 | -98.2056
3.5779 | -98.3919
3.5816 | -98.5096
3.5853 | -98.5285
3.589 | -98.4721
3.5927 | -98.3584
3.5965 | -98.2245
3.6002 | -98.1107
3.6039 | -98.0426
3.6076 | -98.0294
3.6113 | -98.0619
3.6151 | -98.123
3.6188 | -98.1767
3.6225 | -98.1909
3.6262 | -98.1563
3.6299 | -98.0679
3.6337 | -97.9584
3.6374 | -97.8522
3.6411 | -97.7773
3.6448 | -97.7578
3.6485 | -97.8029
3.6523 | -97.8939
3.656 | -97.9991
3.6597 | -98.0869
3.6634 | -98.1464
3.6671 | -98.1686
3.6709 | -98.1539
3.6746 | -98.1265
3.6783 | -98.1047
3.682 | -98.0954
3.6857 | -98.0908
3.6894 | -98.0784
3.6932 | -98.0394
3.6969 | -97.9661
3.7006 | -97.877
3.7043 | -97.7932
3.708 | -97.7328
3.7118 | -97.7003
3.7155 | -97.7003
3.7192 | -97.7233
3.7229 | -97.7656
3.7266 | -97.8308
3.7304 | -97.9125
3.7341 | -98.0053
3.7378 | -98.1064
3.7415 | -98.1963
3.7452 | -98.2613
3.749 | -98.3039
3.7527 | -98.2765
3.7564 | -98.2696
3.7601 | -98.2413
3.7638 | -98.2135
3.7675 | -98.1874
3.7713 | -98.1733
3.775 | -98.1761
3.7787 | -98.1976
3.7824 | -98.2295
3.7861 | -98.2742
3.7899 | -98.3378
3.7936 | -98.4077
3.7973 | -98.4819
3.801 | -98.5533
3.8047 | -98.5936
3.8085 | -98.6012
3.8122 | -98.579
3.8159 | -98.5293
3.8196 | -98.4725
3.8233 | -98.4139
3.8271 | -98.3651
3.8308 | -98.3467
3.8345 | -98.3505
3.8382 | -98.3777
3.8419 | -98.4316
3.8457 | -98.5039
3.8494 | -98.5811
3.8531 | -98.655
3.8568 | -98.7218
3.8605 | -98.7787
3.8642 | -98.8241
3.868 | -98.8707
3.8717 | -98.9205
3.8754 | -98.9627
3.8791 | -98.9881
3.8828 | -98.9918
3.8866 | -98.9779
3.8903 | -98.9522
3.894 | -98.9392
3.8977 | -98.9549
3.9014 | -99.0058
3.9052 | -99.0741
3.9089 | -99.1472
3.9126 | -99.2026
3.9163 | -99.2176
3.92 | -99.1923
3.9238 | -99.1335
3.9275 | -99.0503
3.9312 | -98.9581
3.9349 | -98.8627
3.9386 | -98.796
3.9424 | -98.7702
3.9461 | -98.7802
3.9498 | -98.8262
3.9535 | -98.9045
3.9572 | -98.9899
3.9609 | -99.0694
3.9647 | -99.1222
3.9684 | -99.1546
3.9721 | -99.1566
3.9758 | -99.1345
3.9795 | -99.1079
3.9833 | -99.0874
3.987 | -99.0705
3.9907 | -99.0611
3.9944 | -99.0558
3.9981 | -99.0499
4.0019 | -99.0366
4.0056 | -99.0153
4.0093 | -98.9953
4.013 | -98.9812
4.0167 | -98.9771
4.0205 | -98.9694
4.0242 | -98.9592
4.0279 | -98.9278
4.0316 | -98.8548
4.0353 | -98.7444
4.0391 | -98.6177
4.0428 | -98.4797
4.0465 | -98.348
4.0502 | -98.2437
4.0539 | -98.1816
4.0576 | -98.1729
4.0614 | -98.2245
4.0651 | -98.3317
4.0688 | -98.4861
4.0725 | -98.6722
4.0762 | -98.8675
4.08 | -99.0568
4.0837 | -99.2321
4.0874 | -99.365
4.0911 | -99.4623
4.0948 | -99.5281
4.0986 | -99.566
4.1023 | -99.5753
4.106 | -99.5852
4.1097 | -99.6008
4.1134 | -99.6307
4.1172 | -99.673
4.1209 | -99.7229
4.1246 | -99.7635
4.1283 | -99.777
4.132 | -99.7602
4.1358 | -99.7164
4.1395 | -99.6477
4.1432 | -99.5641
4.1469 | -99.486
4.1506 | -99.4292
4.1543 | -99.3892
4.1581 | -99.3714
4.1618 | -99.3717
4.1655 | -99.3859
4.1692 | -99.3963
4.1729 | -99.3926
4.1767 | -99.3707
4.1804 | -99.3298
4.1841 | -99.2806
4.1878 | -99.2446
4.1915 | -99.2366
4.1953 | -99.2516
4.199 | -99.2816
4.2027 | -99.3164
4.2064 | -99.3573
4.2101 | -99.3959
4.2139 | -99.4316
4.2176 | -99.4626
4.2213 | -99.4861
4.225 | -99.4928
4.2287 | -99.4803
4.2325 | -99.4459
4.2362 | -99.3829
4.2399 | -99.3027
4.2436 | -99.2115
4.2473 | -99.1238
4.251 | -99.0473
4.2548 | -98.9841
4.2585 | -98.9463
4.2622 | -98.9472
4.2659 | -98.9834
4.2696 | -99.042
4.2734 | -99.1359
4.2771 | -99.2564
4.2808 | -99.4027
4.2845 | -99.5604
4.2882 | -99.7146
4.292 | -99.7657
4.2957 | -99.8267
4.2994 | -99.85
4.3031 | -99.8335
4.3068 | -99.7795
4.3106 | -99.7126
4.3143 | -99.6343
4.318 | -99.5582
4.3217 | -99.4944
4.3254 | -99.4424
4.3291 | -99.4045
4.3329 | -99.3814
4.3366 | -99.3746
4.3403 | -99.3698
4.344 | -99.3645
4.3477 | -99.3663
4.3515 | -99.3661
4.3552 | -99.3741
4.3589 | -99.3914
4.3626 | -99.4023
4.3663 | -99.4142
4.3701 | -99.4199
4.3738 | -99.423
4.3775 | -99.4146
4.3812 | -99.3976
4.3849 | -99.3691
4.3887 | -99.3308
4.3924 | -99.2932
4.3961 | -99.2616
4.3998 | -99.2398
4.4035 | -99.2477
4.4073 | -99.283
4.411 | -99.3492
4.4147 | -99.4399
4.4184 | -99.5393
4.4221 | -99.6376
4.4258 | -99.7292
4.4296 | -99.802
4.4333 | -99.8489
4.437 | -99.856
4.4407 | -99.8389
4.4444 | -99.7918
4.4482 | -99.7306
4.4519 | -99.673
4.4556 | -99.6354
4.4593 | -99.612
4.463 | -99.6057
4.4668 | -99.6157
4.4705 | -99.6314
4.4742 | -99.6462
4.4779 | -99.6599
4.4816 | -99.6592
4.4854 | -99.6468
4.4891 | -99.6221
4.4928 | -99.5895
4.4965 | -99.5472
4.5002 | -99.5155
4.504 | -99.5029
4.5077 | -99.5039
4.5114 | -99.5239
4.5151 | -99.5591
4.5188 | -99.593
4.5225 | -99.6092
4.5263 | -99.6101
4.53 | -99.5985
4.5337 | -99.573
4.5374 | -99.5557
4.5411 | -99.5492
4.5449 | -99.5463
4.5486 | -99.5521
4.5523 | -99.5674
4.556 | -99.5837
4.5597 | -99.6071
4.5635 | -99.6408
4.5672 | -99.6791
4.5709 | -99.7218
4.5746 | -99.7615
4.5783 | -99.7831
4.5821 | -99.7786
4.5858 | -99.7498
4.5895 | -99.709
4.5932 | -99.6737
4.5969 | -99.6478
4.6007 | -99.6475
4.6044 | -99.6682
4.6081 | -99.6999
4.6118 | -99.7575
4.6155 | -99.8386
4.6192 | -99.9295
4.623 | -100.0284
4.6267 | -100.1294
4.6304 | -100.2205
4.6341 | -100.3023
4.6378 | -100.3711
4.6416 | -100.4138
4.6453 | -100.4309
4.649 | -100.4256
4.6527 | -100.4012
4.6564 | -100.3666
4.6602 | -100.3166
4.6639 | -100.2649
4.6676 | -100.2363
4.6713 | -100.2218
4.675 | -100.2266
4.6788 | -100.2444
4.6825 | -100.2517
4.6862 | -100.2555
4.6899 | -100.2444
4.6936 | -100.2154
4.6974 | -100.1621
4.7011 | -100.0895
4.7048 | -99.9907
4.7085 | -99.8803
4.7122 | -99.7753
4.7159 | -99.6868
4.7197 | -99.6266
4.7234 | -99.5979
4.7271 | -99.6024
4.7308 | -99.6385
4.7345 | -99.6948
4.7383 | -99.7535
4.742 | -99.8102
4.7457 | -99.8552
4.7494 | -99.8693
4.7531 | -99.8654
4.7569 | -99.8453
4.7606 | -99.8212
4.7643 | -99.8001
4.768 | -99.7834
4.7717 | -99.7778
4.7755 | -99.8001
4.7792 | -99.8336
4.7829 | -99.8746
4.7866 | -99.9256
4.7903 | -99.9815
4.794 | -100.0357
4.7978 | -100.0913
4.8015 | -100.1463
4.8052 | -100.1965
4.8089 | -100.256
4.8126 | -100.3129
4.8164 | -100.363
4.8201 | -100.4063
4.8238 | -100.4241
4.8275 | -100.3651
4.8312 | -100.3281
4.835 | -100.2716
4.8387 | -100.2036
4.8424 | -100.1369
4.8461 | -100.0687
4.8498 | -100.0092
4.8536 | -99.9731
4.8573 | -99.96
4.861 | -99.9701
4.8647 | -100.0049
4.8684 | -100.0692
4.8722 | -100.1557
4.8759 | -100.2561
4.8796 | -100.3664
4.8833 | -100.4774
4.887 | -100.5964
4.8907 | -100.7187
4.8945 | -100.8359
4.8982 | -100.94
4.9019 | -101.0245
4.9056 | -101.0804
4.9093 | -101.1075
4.9131 | -101.1094
4.9168 | -101.0776
4.9205 | -101.025
4.9242 | -100.9597
4.9279 | -100.875
4.9317 | -100.7791
4.9354 | -100.6758
4.9391 | -100.5684
4.9428 | -100.4657
4.9465 | -100.3649
4.9503 | -100.2711
4.954 | -100.1984
4.9577 | -100.1473
4.9614 | -100.0832
4.9651 | -100.029
4.9689 | -99.9921
4.9726 | -99.9573
4.9763 | -99.9143
4.98 | -99.8582
4.9837 | -99.7855
4.9874 | -99.7
4.9912 | -99.6129
4.9949 | -99.5371
4.9986 | -99.4861
5.0023 | -99.475
5.006 | -99.5128
5.0098 | -99.5988
5.0135 | -99.723
5.0172 | -99.8713
5.0209 | -100.0376
5.0246 | -100.2092
5.0284 | -100.3679
5.0321 | -100.4964
5.0358 | -100.5911
5.0395 | -100.6782
5.0432 | -100.7226
5.047 | -100.7215
5.0507 | -100.7005
5.0544 | -100.6728
5.0581 | -100.6411
5.0618 | -100.6059
5.0656 | -100.5757
5.0693 | -100.5449
5.073 | -100.5154
5.0767 | -100.4859
5.0804 | -100.4578
5.0841 | -100.4288
5.0879 | -100.3953
5.0916 | -100.3622
5.0953 | -100.3282
5.099 | -100.2888
5.1027 | -100.2521
5.1065 | -100.2366
5.1102 | -100.2439
5.1139 | -100.2715
5.1176 | -100.3205
5.1213 | -100.3851
5.1251 | -100.4588
5.1288 | -100.5393
5.1325 | -100.6063
5.1362 | -100.6557
5.1399 | -100.6939
5.1437 | -100.717
5.1474 | -100.7318
5.1511 | -100.7487
5.1548 | -100.7656
5.1585 | -100.7833
5.1623 | -100.8015
5.166 | -100.824
5.1697 | -100.848
5.1734 | -100.8726
5.1771 | -100.8978
5.1808 | -100.9241
5.1846 | -100.9448
5.1883 | -100.9704
5.192 | -101.0021
5.1957 | -101.0347
5.1994 | -101.072
5.2032 | -101.1024
5.2069 | -101.1191
5.2106 | -101.1339
5.2143 | -101.1393
5.218 | -101.1428
5.2218 | -101.1503
5.2255 | -101.1525
5.2292 | -101.143
5.2329 | -101.1282
5.2366 | -101.1076
5.2404 | -101.0869
5.2441 | -101.0649
5.2478 | -101.0434
5.2515 | -101.0273
5.2552 | -101.0131
5.2589 | -101.0008
5.2627 | -100.9868
5.2664 | -100.9678
5.2701 | -100.9418
5.2738 | -100.907
5.2775 | -100.8675
5.2813 | -100.8426
5.285 | -100.8287
5.2887 | -100.8166
5.2924 | -100.8189
5.2961 | -100.8281
5.2999 | -100.8429
5.3036 | -100.8695
5.3073 | -100.896
5.311 | -100.9242
5.3147 | -100.9507
5.3185 | -100.9752
5.3222 | -101.0002
5.3259 | -101.0255
5.3296 | -101.0491
5.3333 | -101.0745
5.3371 | -101.1045
5.3408 | -101.1333
5.3445 | -101.1576
5.3482 | -101.1843
5.3519 | -101.218
5.3556 | -101.2492
5.3594 | -101.2715
5.3631 | -101.2852
5.3668 | -101.299
5.3705 | -101.2937
5.3742 | -101.2581
5.378 | -101.2049
5.3817 | -101.1418
5.3854 | -101.0585
5.3891 | -100.9645
5.3928 | -100.8602
5.3966 | -100.7499
5.4003 | -100.6476
5.404 | -100.5527
5.4077 | -100.4745
5.4114 | -100.4142
5.4152 | -100.371
5.4189 | -100.3435
5.4226 | -100.3347
5.4263 | -100.3433
5.43 | -100.3609
5.4338 | -100.3908
5.4375 | -100.4346
5.4412 | -100.4878
5.4449 | -100.5442
5.4486 | -100.61
5.4523 | -100.6799
5.4561 | -100.7664
5.4598 | -100.8582
5.4635 | -100.9375
5.4672 | -101.0093
5.4709 | -101.086
5.4747 | -101.151
5.4784 | -101.197
5.4821 | -101.2367
5.4858 | -101.2504
5.4895 | -101.2471
5.4933 | -101.2302
5.497 | -101.2061
5.5007 | -101.176
5.5044 | -101.1395
5.5081 | -101.1099
5.5119 | -101.0901
5.5156 | -101.0887
5.5193 | -101.0899
5.523 | -101.1029
5.5267 | -101.1344
5.5305 | -101.184
5.5342 | -101.2506
5.5379 | -101.3302
5.5416 | -101.4083
5.5453 | -101.4785
5.549 | -101.5254
5.5528 | -101.5535
5.5565 | -101.5541
5.5602 | -101.5424
5.5639 | -101.5203
5.5676 | -101.4856
5.5714 | -101.4542
5.5751 | -101.4219
5.5788 | -101.3909
5.5825 | -101.3692
5.5862 | -101.3612
5.59 | -101.3652
5.5937 | -101.3718
5.5974 | -101.373
5.6011 | -101.3788
5.6048 | -101.3916
5.6086 | -101.3989
5.6123 | -101.4046
5.616 | -101.4127
5.6197 | -101.4247
5.6234 | -101.4397
5.6272 | -101.4624
5.6309 | -101.4942
5.6346 | -101.5416
5.6383 | -101.5862
5.642 | -101.6293
5.6457 | -101.6634
5.6495 | -101.6929
5.6532 | -101.7147
5.6569 | -101.7301
5.6606 | -101.7373
5.6643 | -101.7378
5.6681 | -101.722
5.6718 | -101.6997
5.6755 | -101.6742
5.6792 | -101.6485
5.6829 | -101.6204
5.6867 | -101.5808
5.6904 | -101.5494
5.6941 | -101.5204
5.6978 | -101.4805
5.7015 | -101.4374
5.7053 | -101.3958
5.709 | -101.363
5.7127 | -101.3306
5.7164 | -101.3069
5.7201 | -101.2977
5.7238 | -101.3016
5.7276 | -101.322
5.7313 | -101.3636
5.735 | -101.4142
5.7387 | -101.4676
5.7424 | -101.5171
5.7462 | -101.5643
5.7499 | -101.5942
5.7536 | -101.6151
5.7573 | -101.6181
5.761 | -101.6074
5.7648 | -101.5866
5.7685 | -101.5658
5.7722 | -101.5454
5.7759 | -101.5244
5.7796 | -101.5212
5.7834 | -101.5339
5.7871 | -101.5582
5.7908 | -101.5907
5.7945 | -101.6305
5.7982 | -101.6765
5.802 | -101.7152
5.8057 | -101.7425
5.8094 | -101.7613
5.8131 | -101.7661
5.8168 | -101.756
5.8205 | -101.7377
5.8243 | -101.7073
5.828 | -101.6748
5.8317 | -101.6318
5.8354 | -101.5876
5.8391 | -101.5369
5.8429 | -101.483
5.8466 | -101.4333
5.8503 | -101.3799
5.854 | -101.3387
5.8577 | -101.3119
5.8615 | -101.3026
5.8652 | -101.3155
5.8689 | -101.3551
5.8726 | -101.4223
5.8763 | -101.5183
5.8801 | -101.6386
5.8838 | -101.7738
5.8875 | -101.9147
5.8912 | -102.0527
5.8949 | -102.1777
5.8987 | -102.219
5.9024 | -102.2907
5.9061 | -102.3347
5.9098 | -102.347
5.9135 | -102.3415
5.9172 | -102.309
5.921 | -102.263
5.9247 | -102.2063
5.9284 | -102.1501
5.9321 | -102.1017
5.9358 | -102.0632
5.9396 | -102.0438
5.9433 | -102.0283
5.947 | -102.0193
5.9507 | -102.0184
5.9544 | -102.0256
5.9582 | -102.04
5.9619 | -102.0501
5.9656 | -102.0512
5.9693 | -102.0459
5.973 | -102.0346
5.9768 | -102.0209
5.9805 | -102.0041
5.9842 | -101.9822
5.9879 | -101.9525
5.9916 | -101.9213
5.9954 | -101.8905
5.9991 | -101.8582
6.0028 | -101.8293
6.0065 | -101.8045
6.0102 | -101.7809
6.0139 | -101.7549
6.0177 | -101.73
6.0214 | -101.7026
6.0251 | -101.6717
6.0288 | -101.6411
6.0325 | -101.6124
6.0363 | -101.589
6.04 | -101.5717
6.0437 | -101.5621
6.0474 | -101.5706
6.0511 | -101.5915
6.0549 | -101.6299
6.0586 | -101.6846
6.0623 | -101.7421
6.066 | -101.8036
6.0697 | -101.8645
6.0735 | -101.9181
6.0772 | -101.9607
6.0809 | -101.9963
6.0846 | -102.0155
6.0883 | -102.0217
6.0921 | -102.0181
6.0958 | -102.0052
6.0995 | -101.9753
6.1032 | -101.9457
6.1069 | -101.9103
6.1106 | -101.8616
6.1144 | -101.8135
6.1181 | -101.7662
6.1218 | -101.7194
6.1255 | -101.6764
6.1292 | -101.6414
6.133 | -101.6209
6.1367 | -101.6114
6.1404 | -101.6105
6.1441 | -101.6269
6.1478 | -101.658
6.1516 | -101.69
6.1553 | -101.7283
6.159 | -101.764
6.1627 | -101.7849
6.1664 | -101.8005
6.1702 | -101.8129
6.1739 | -101.81
6.1776 | -101.8011
6.1813 | -101.79
6.185 | -101.7791
6.1887 | -101.7653
6.1925 | -101.7613
6.1962 | -101.7618
6.1999 | -101.7752
6.2036 | -101.8016
6.2073 | -101.8286
6.2111 | -101.8621
6.2148 | -101.8935
6.2185 | -101.9208
6.2222 | -101.9469
6.2259 | -101.9605
6.2297 | -101.9692
6.2334 | -101.9677
6.2371 | -101.96
6.2408 | -101.946
6.2445 | -101.9374
6.2483 | -101.935
6.252 | -101.9393
6.2557 | -101.9555
6.2594 | -101.9816
6.2631 | -102.0227
6.2669 | -102.0885
6.2706 | -102.1657
6.2743 | -102.2389
6.278 | -102.3066
6.2817 | -102.3688
6.2854 | -102.414
6.2892 | -102.4382
6.2929 | -102.4392
6.2966 | -102.4169
6.3003 | -102.3795
6.304 | -102.3286
6.3078 | -102.2713
6.3115 | -102.2085
6.3152 | -102.137
6.3189 | -102.0674
6.3226 | -102.0051
6.3264 | -101.9543
6.3301 | -101.9149
6.3338 | -101.8885
6.3375 | -101.8763
6.3412 | -101.8803
6.345 | -101.9013
6.3487 | -101.9395
6.3524 | -101.9909
6.3561 | -102.057
6.3598 | -102.128
6.3636 | -102.2059
6.3673 | -102.2931
6.371 | -102.3741
6.3747 | -102.4426
6.3784 | -102.4965
6.3821 | -102.5427
6.3859 | -102.5771
6.3896 | -102.6046
6.3933 | -102.6155
6.397 | -102.6145
6.4007 | -102.6071
6.4045 | -102.5925
6.4082 | -102.5747
6.4119 | -102.5546
6.4156 | -102.5349
6.4193 | -102.5211
6.4231 | -102.5128
6.4268 | -102.5081
6.4305 | -102.5048
6.4342 | -102.5012
6.4379 | -102.5088
6.4417 | -102.5081
6.4454 | -102.5151
6.4491 | -102.5251
6.4528 | -102.535
6.4565 | -102.5512
6.4603 | -102.5711
6.464 | -102.5929
6.4677 | -102.6191
6.4714 | -102.6482
6.4751 | -102.6764
6.4788 | -102.6984
6.4826 | -102.7165
6.4863 | -102.7304
6.49 | -102.7413
6.4937 | -102.7418
6.4974 | -102.7302
6.5012 | -102.7159
6.5049 | -102.6918
6.5086 | -102.6554
6.5123 | -102.6071
6.516 | -102.5499
6.5198 | -102.4916
6.5235 | -102.4335
6.5272 | -102.3793
6.5309 | -102.323
6.5346 | -102.2746
6.5384 | -102.2385
6.5421 | -102.2106
6.5458 | -102.1873
6.5495 | -102.1709
6.5532 | -102.1621
6.5569 | -102.1594
6.5607 | -102.1645
6.5644 | -102.1716
6.5681 | -102.1799
6.5718 | -102.192
6.5755 | -102.2017
6.5793 | -102.2161
6.583 | -102.2348
6.5867 | -102.2526
6.5904 | -102.2693
6.5941 | -102.2883
6.5979 | -102.3081
6.6016 | -102.3195
6.6053 | -102.3332
6.609 | -102.3514
6.6127 | -102.3667
6.6165 | -102.3833
6.6202 | -102.4033
6.6239 | -102.4239
6.6276 | -102.4438
6.6313 | -102.4677
6.6351 | -102.482
6.6388 | -102.488
6.6425 | -102.495
6.6462 | -102.4969
6.6499 | -102.4926
6.6536 | -102.4837
6.6574 | -102.4707
6.6611 | -102.4572
6.6648 | -102.445
6.6685 | -102.4268
6.6722 | -102.4022
6.676 | -102.3868
6.6797 | -102.3719
6.6834 | -102.359
6.6871 | -102.3537
6.6908 | -102.3577
6.6946 | -102.3683
6.6983 | -102.384
6.702 | -102.4109
6.7057 | -102.4432
6.7094 | -102.4736
6.7132 | -102.508
6.7169 | -102.5353
6.7206 | -102.5533
6.7243 | -102.5699
6.728 | -102.5788
6.7318 | -102.583
6.7355 | -102.5899
6.7392 | -102.5951
6.7429 | -102.5938
6.7466 | -102.5926
6.7503 | -102.5964
6.7541 | -102.6064
6.7578 | -102.622
6.7615 | -102.6431
6.7652 | -102.6716
6.7689 | -102.7091
6.7727 | -102.7547
6.7764 | -102.7962
6.7801 | -102.8439
6.7838 | -102.8947
6.7875 | -102.943
6.7913 | -102.9879
6.795 | -103.0274
6.7987 | -103.0667
6.8024 | -103.0942
6.8061 | -103.1134
6.8099 | -103.1266
6.8136 | -103.1235
6.8173 | -103.1126
6.821 | -103.0908
6.8247 | -103.0615
6.8285 | -103.0278
6.8322 | -102.9854
6.8359 | -102.9408
6.8396 | -102.8974
6.8433 | -102.859
6.847 | -102.8259
6.8508 | -102.7971
6.8545 | -102.7679
6.8582 | -102.7516
6.8619 | -102.7422
6.8656 | -102.7421
6.8694 | -102.7515
6.8731 | -102.7676
6.8768 | -102.7894
6.8805 | -102.8155
6.8842 | -102.8473
6.888 | -102.879
6.8917 | -102.9102
6.8954 | -102.9408
6.8991 | -102.9598
6.9028 | -102.9742
6.9066 | -102.9803
6.9103 | -102.9772
6.914 | -102.9698
6.9177 | -102.9547
6.9214 | -102.9297
6.9252 | -102.9042
6.9289 | -102.8804
6.9326 | -102.8635
6.9363 | -102.8518
6.94 | -102.8487
6.9437 | -102.8432
6.9475 | -102.8398
6.9512 | -102.8225
6.9549 | -102.8106
6.9586 | -102.7935
6.9623 | -102.7677
6.9661 | -102.7299
6.9698 | -102.6793
6.9735 | -102.6001
6.9772 | -102.5437
6.9809 | -102.4849
6.9847 | -102.4274
6.9884 | -102.3736
6.9921 | -102.3333
6.9958 | -102.3048
6.9995 | -102.2882
7.0033 | -102.29
7.007 | -102.312
7.0107 | -102.3521
7.0144 | -102.4092
7.0181 | -102.4787
7.0218 | -102.5558
7.0256 | -102.6357
7.0293 | -102.7078
7.033 | -102.7715
7.0367 | -102.8216
7.0404 | -102.8495
7.0442 | -102.8577
7.0479 | -102.8489
7.0516 | -102.822
7.0553 | -102.7895
7.059 | -102.7344
7.0628 | -102.672
7.0665 | -102.6076
7.0702 | -102.549
7.0739 | -102.4976
7.0776 | -102.4434
7.0814 | -102.4029
7.0851 | -102.3742
7.0888 | -102.358
7.0925 | -102.349
7.0962 | -102.3515
7.1 | -102.3645
7.1037 | -102.3855
7.1074 | -102.4086
7.1111 | -102.4421
7.1148 | -102.4819
7.1185 | -102.5227
7.1223 | -102.5667
7.126 | -102.6096
7.1297 | -102.6551
7.1334 | -102.7014
7.1371 | -102.7431
7.1409 | -102.7825
7.1446 | -102.8226
7.1483 | -102.864
7.152 | -102.9041
7.1557 | -102.9435
7.1595 | -102.9749
7.1632 | -102.9962
7.1669 | -103.0184
7.1706 | -103.0351
7.1743 | -103.0548
7.1781 | -103.0665
7.1818 | -103.0758
7.1855 | -103.0932
7.1892 | -103.1016
7.1929 | -103.1063
7.1967 | -103.117
7.2004 | -103.1309
7.2041 | -103.145
7.2078 | -103.1609
7.2115 | -103.1776
7.2152 | -103.2032
7.219 | -103.2219
7.2227 | -103.2396
7.2264 | -103.2611
7.2301 | -103.2818
7.2338 | -103.3014
7.2376 | -103.3116
7.2413 | -103.3156
7.245 | -103.3146
7.2487 | -103.3038
7.2524 | -103.2762
7.2562 | -103.2467
7.2599 | -103.2166
7.2636 | -103.1874
7.2673 | -103.165
7.271 | -103.1528
7.2748 | -103.1393
7.2785 | -103.1282
7.2822 | -103.12
7.2859 | -103.1203
7.2896 | -103.1265
7.2934 | -103.1382
7.2971 | -103.1543
7.3008 | -103.1766
7.3045 | -103.1958
7.3082 | -103.2094
7.3119 | -103.2236
7.3157 | -103.234
7.3194 | -103.2408
7.3231 | -103.2356
7.3268 | -103.2262
7.3305 | -103.2108
7.3343 | -103.1913
7.338 | -103.1697
7.3417 | -103.1453
7.3454 | -103.1285
7.3491 | -103.1151
7.3529 | -103.1029
7.3566 | -103.0986
7.3603 | -103.1034
7.364 | -103.1084
7.3677 | -103.1163
7.3715 | -103.1307
7.3752 | -103.1511
7.3789 | -103.1708
7.3826 | -103.1882
7.3863 | -103.2085
7.3901 | -103.2299
7.3938 | -103.2427
7.3975 | -103.246
7.4012 | -103.2407
7.4049 | -103.2296
7.4086 | -103.2086
7.4124 | -103.1848
7.4161 | -103.1621
7.4198 | -103.1292
7.4235 | -103.0935
7.4272 | -103.0598
7.431 | -103.0269
7.4347 | -102.9906
7.4384 | -102.9519
7.4421 | -102.915
7.4458 | -102.8841
7.4496 | -102.8623
7.4533 | -102.842
7.457 | -102.8267
7.4607 | -102.8211
7.4644 | -102.8206
7.4682 | -102.8235
7.4719 | -102.8263
7.4756 | -102.8288
7.4793 | -102.8299
7.483 | -102.8301
7.4867 | -102.8323
7.4905 | -102.836
7.4942 | -102.8352
7.4979 | -102.8359
7.5016 | -102.8415
7.5053 | -102.8546
7.5091 | -102.8923
7.5128 | -102.9176
7.5165 | -102.9511
7.5202 | -102.9928
7.5239 | -103.0389
7.5277 | -103.0905
7.5314 | -103.1493
7.5351 | -103.2084
7.5388 | -103.2668
7.5425 | -103.3247
7.5463 | -103.3836
7.55 | -103.436
7.5537 | -103.48
7.5574 | -103.5104
7.5611 | -103.5265
7.5649 | -103.5369
7.5686 | -103.5406
7.5723 | -103.5347
7.576 | -103.5199
7.5797 | -103.5043
7.5835 | -103.4856
7.5872 | -103.4675
7.5909 | -103.4504
7.5946 | -103.4346
7.5983 | -103.4243
7.602 | -103.4187
7.6058 | -103.4212
7.6095 | -103.4268
7.6132 | -103.4365
7.6169 | -103.4456
7.6206 | -103.4601
7.6244 | -103.4763
7.6281 | -103.4959
7.6318 | -103.5115
7.6355 | -103.5207
7.6392 | -103.5253
7.643 | -103.5228
7.6467 | -103.5147
7.6504 | -103.5083
7.6541 | -103.5015
7.6578 | -103.4899
7.6616 | -103.4715
7.6653 | -103.4494
7.669 | -103.4215
7.6727 | -103.3939
7.6764 | -103.3681
7.6801 | -103.3409
7.6839 | -103.318
7.6876 | -103.3002
7.6913 | -103.2909
7.695 | -103.292
7.6987 | -103.3052
7.7025 | -103.3237
7.7062 | -103.3471
7.7099 | -103.3746
7.7136 | -103.4029
7.7173 | -103.4326
7.7211 | -103.4597
7.7248 | -103.4855
7.7285 | -103.5099
7.7322 | -103.5294
7.7359 | -103.5442
7.7397 | -103.5539
7.7434 | -103.5523
7.7471 | -103.5515
7.7508 | -103.5481
7.7545 | -103.5444
7.7583 | -103.542
7.762 | -103.5352
7.7657 | -103.5228
7.7694 | -103.5062
7.7731 | -103.4955
7.7768 | -103.4926
7.7806 | -103.496
7.7843 | -103.5063
7.788 | -103.5142
7.7917 | -103.5229
7.7954 | -103.5347
7.7992 | -103.544
7.8029 | -103.5555
7.8066 | -103.5603
7.8103 | -103.5606
7.814 | -103.5606
7.8178 | -103.5584
7.8215 | -103.5553
7.8252 | -103.5546
7.8289 | -103.5543
7.8326 | -103.5504
7.8364 | -103.5481
7.8401 | -103.5451
7.8438 | -103.5541
7.8475 | -103.5705
7.8512 | -103.5892
7.855 | -103.6176
7.8587 | -103.656
7.8624 | -103.6885
7.8661 | -103.7219
7.8698 | -103.7527
7.8735 | -103.7798
7.8773 | -103.8051
7.881 | -103.8225
7.8847 | -103.8349
7.8884 | -103.8412
7.8921 | -103.8452
7.8959 | -103.8485
7.8996 | -103.8503
7.9033 | -103.8601
7.907 | -103.8719
7.9107 | -103.8805
7.9145 | -103.8908
7.9182 | -103.8976
7.9219 | -103.9083
7.9256 | -103.9184
7.9293 | -103.9282
7.9331 | -103.9339
7.9368 | -103.9342
7.9405 | -103.9282
7.9442 | -103.9218
7.9479 | -103.9167
7.9516 | -103.904
7.9554 | -103.8845
7.9591 | -103.8504
7.9628 | -103.8044
7.9665 | -103.7564
7.9702 | -103.7073
7.974 | -103.6535
7.9777 | -103.6024
7.9814 | -103.557
7.9851 | -103.5135
7.9888 | -103.4709
7.9926 | -103.4371
7.9963 | -103.4095
8 | -103.3887
8.0037 | -103.3714
8.0074 | -103.3548
8.0112 | -103.348
8.0149 | -103.3513
8.0186 | -103.3647
8.0223 | -103.3845
8.026 | -103.4109
8.0298 | -103.4413
8.0335 | -103.4786
8.0372 | -103.5153
8.0409 | -103.5488
8.0446 | -103.5818
8.0483 | -103.6132
8.0521 | -103.6484
8.0558 | -103.6851
8.0595 | -103.7183
8.0632 | -103.751
8.0669 | -103.7821
8.0707 | -103.8167
8.0744 | -103.855
8.0781 | -103.8888
8.0818 | -103.9211
8.0855 | -103.9511
8.0893 | -103.9734
8.093 | -103.9933
8.0967 | -104.0057
8.1004 | -104.0062
8.1041 | -104.006
8.1079 | -104.0074
8.1116 | -103.9988
8.1153 | -103.9901
8.119 | -103.9743
8.1227 | -103.9456
8.1265 | -103.9182
8.1302 | -103.8835
8.1339 | -103.8473
8.1376 | -103.8071
8.1413 | -103.7736
8.145 | -103.7402
8.1488 | -103.7063
8.1525 | -103.6654
8.1562 | -103.6215
8.1599 | -103.581
8.1636 | -103.5446
8.1674 | -103.5099
8.1711 | -103.4803
8.1748 | -103.4542
8.1785 | -103.4329
8.1822 | -103.4148
8.186 | -103.407
8.1897 | -103.4059
8.1934 | -103.4102
8.1971 | -103.4251
8.2008 | -103.4493
8.2046 | -103.4809
8.2083 | -103.514
8.212 | -103.5488
8.2157 | -103.5871
8.2194 | -103.6291
8.2232 | -103.6647
8.2269 | -103.6956
8.2306 | -103.7246
8.2343 | -103.7558
8.238 | -103.7828
8.2417 | -103.8084
8.2455 | -103.8356
8.2492 | -103.855
8.2529 | -103.8753
8.2566 | -103.8904
8.2603 | -103.9023
8.2641 | -103.9074
8.2678 | -103.9124
8.2715 | -103.9153
8.2752 | -103.9162
8.2789 | -103.9164
8.2827 | -103.908
8.2864 | -103.8929
8.2901 | -103.8769
8.2938 | -103.8581
8.2975 | -103.8353
8.3013 | -103.815
8.305 | -103.8034
8.3087 | -103.7892
8.3124 | -103.7762
8.3161 | -103.767
8.3199 | -103.7557
8.3236 | -103.747
8.3273 | -103.741
8.331 | -103.7372
8.3347 | -103.7362
8.3384 | -103.7378
8.3422 | -103.7424
8.3459 | -103.7546
8.3496 | -103.7631
8.3533 | -103.7698
8.357 | -103.7772
8.3608 | -103.7814
8.3645 | -103.7866
8.3682 | -103.795
8.3719 | -103.8042
8.3756 | -103.8113
8.3794 | -103.8205
8.3831 | -103.829
8.3868 | -103.839
8.3905 | -103.8533
8.3942 | -103.8671
8.398 | -103.8828
8.4017 | -103.8997
8.4054 | -103.9168
8.4091 | -103.9302
8.4128 | -103.9422
8.4166 | -103.9507
8.4203 | -103.9624
8.424 | -103.972
8.4277 | -103.9735
8.4314 | -103.9741
8.4351 | -103.9746
8.4389 | -103.9755
8.4426 | -103.9778
8.4463 | -103.9779
8.45 | -103.9779
8.4537 | -103.9762
8.4575 | -103.9754
8.4612 | -103.9808
8.4649 | -103.9881
8.4686 | -103.9962
8.4723 | -104.0112
8.4761 | -104.0301
8.4798 | -104.0485
8.4835 | -104.0709
8.4872 | -104.0953
8.4909 | -104.1169
8.4947 | -104.1373
8.4984 | -104.1515
8.5021 | -104.1673
8.5058 | -104.1804
8.5095 | -104.1898
8.5132 | -104.1945
8.517 | -104.1985
8.5207 | -104.2008
8.5244 | -104.1988
8.5281 | -104.1864
8.5318 | -104.1725
8.5356 | -104.1539
8.5393 | -104.1385
8.543 | -104.1214
8.5467 | -104.0991
8.5504 | -104.082
8.5542 | -104.0655
8.5579 | -104.0525
8.5616 | -104.0455
8.5653 | -104.0447
8.569 | -104.049
8.5728 | -104.0563
8.5765 | -104.0697
8.5802 | -104.09
8.5839 | -104.1156
8.5876 | -104.1286
8.5914 | -104.1414
8.5951 | -104.155
8.5988 | -104.167
8.6025 | -104.1745
8.6062 | -104.1809
8.6099 | -104.183
8.6137 | -104.1765
8.6174 | -104.1637
8.6211 | -104.149
8.6248 | -104.1358
8.6285 | -104.1192
8.6323 | -104.1056
8.636 | -104.0867
8.6397 | -104.0685
8.6434 | -104.0548
8.6471 | -104.0456
8.6509 | -104.0417
8.6546 | -104.0478
8.6583 | -104.0574
8.662 | -104.0734
8.6657 | -104.0877
8.6695 | -104.1064
8.6732 | -104.1344
8.6769 | -104.166
8.6806 | -104.1949
8.6843 | -104.2223
8.6881 | -104.2468
8.6918 | -104.2686
8.6955 | -104.2884
8.6992 | -104.3024
8.7029 | -104.309
8.7066 | -104.308
8.7104 | -104.3024
8.7141 | -104.2931
8.7178 | -104.2811
8.7215 | -104.2709
8.7252 | -104.2593
8.729 | -104.2479
8.7327 | -104.2347
8.7364 | -104.2228
8.7401 | -104.2162
8.7438 | -104.2128
8.7476 | -104.2101
8.7513 | -104.204
8.755 | -104.2007
8.7587 | -104.1934
8.7624 | -104.1859
8.7662 | -104.1854
8.7699 | -104.1886
8.7736 | -104.1917
8.7773 | -104.196
8.781 | -104.2034
8.7848 | -104.2087
8.7885 | -104.2158
8.7922 | -104.2246
8.7959 | -104.2365
8.7996 | -104.2457
8.8033 | -104.2514
8.8071 | -104.2558
8.8108 | -104.2632
8.8145 | -104.2694
8.8182 | -104.2713
8.8219 | -104.2712
8.8257 | -104.2722
8.8294 | -104.2794
8.8331 | -104.2854
8.8368 | -104.2851
8.8405 | -104.2831
8.8443 | -104.2796
8.848 | -104.275
8.8517 | -104.2609
8.8554 | -104.2459
8.8591 | -104.2243
8.8629 | -104.2051
8.8666 | -104.1881
8.8703 | -104.1686
8.874 | -104.1496
8.8777 | -104.1303
8.8814 | -104.1115
8.8852 | -104.0966
8.8889 | -104.0883
8.8926 | -104.0834
8.8963 | -104.0759
8.9 | -104.0685
8.9038 | -104.067
8.9075 | -104.0703
8.9112 | -104.0725
8.9149 | -104.0847
8.9186 | -104.0942
8.9224 | -104.0996
8.9261 | -104.106
8.9298 | -104.1106
8.9335 | -104.1155
8.9372 | -104.1148
8.941 | -104.1166
8.9447 | -104.1159
8.9484 | -104.1098
8.9521 | -104.1011
8.9558 | -104.0853
8.9596 | -104.0599
8.9633 | -104.036
8.967 | -104.016
8.9707 | -103.9902
8.9744 | -103.9603
8.9781 | -103.9333
8.9819 | -103.9148
8.9856 | -103.8966
8.9893 | -103.8835
8.993 | -103.8757
8.9967 | -103.8683
9.0005 | -103.8688
9.0042 | -103.8761
9.0079 | -103.8886
9.0116 | -103.9078
9.0153 | -103.9293
9.0191 | -103.9486
9.0228 | -103.9692
9.0265 | -103.9937
9.0302 | -104.0173
9.0339 | -104.0392
9.0377 | -104.0593
9.0414 | -104.0771
9.0451 | -104.0902
9.0488 | -104.1066
9.0525 | -104.1193
9.0563 | -104.1259
9.06 | -104.1351
9.0637 | -104.1452
9.0674 | -104.1556
9.0711 | -104.1635
9.0748 | -104.1783
9.0786 | -104.2006
9.0823 | -104.2259
9.086 | -104.2584
9.0897 | -104.292
9.0934 | -104.3303
9.0972 | -104.3699
9.1009 | -104.4122
9.1046 | -104.4535
9.1083 | -104.4888
9.112 | -104.5093
9.1158 | -104.5275
9.1195 | -104.5173
9.1232 | -104.5239
9.1269 | -104.5218
9.1306 | -104.5153
9.1344 | -104.4997
9.1381 | -104.4765
9.1418 | -104.4488
9.1455 | -104.4143
9.1492 | -104.3695
9.153 | -104.3224
9.1567 | -104.2734
9.1604 | -104.224
9.1641 | -104.1783
9.1678 | -104.1362
9.1715 | -104.095
9.1753 | -104.0598
9.179 | -104.0282
9.1827 | -104.0021
9.1864 | -103.985
9.1901 | -103.9705
9.1939 | -103.9588
9.1976 | -103.9549
9.2013 | -103.9601
9.205 | -103.9648
9.2087 | -103.968
9.2125 | -103.9702
9.2162 | -103.9746
9.2199 | -103.9841
9.2236 | -103.9973
9.2273 | -104.0037
9.2311 | -104.0124
9.2348 | -104.0226
9.2385 | -104.0325
9.2422 | -104.0433
9.2459 | -104.0593
9.2497 | -104.0725
9.2534 | -104.0816
9.2571 | -104.0921
9.2608 | -104.1046
9.2645 | -104.1204
9.2682 | -104.1376
9.272 | -104.1581
9.2757 | -104.1803
9.2794 | -104.2032
9.2831 | -104.2257
9.2868 | -104.254
9.2906 | -104.2844
9.2943 | -104.3123
9.298 | -104.3334
9.3017 | -104.3478
9.3054 | -104.3578
9.3092 | -104.3735
9.3129 | -104.3858
9.3166 | -104.3914
9.3203 | -104.3917
9.324 | -104.3903
9.3278 | -104.3907
9.3315 | -104.3871
9.3352 | -104.3798
9.3389 | -104.3713
9.3426 | -104.3648
9.3464 | -104.3609
9.3501 | -104.3617
9.3538 | -104.3602
9.3575 | -104.3504
9.3612 | -104.3441
9.3649 | -104.3396
9.3687 | -104.3331
9.3724 | -104.3287
9.3761 | -104.3274
9.3798 | -104.3289
9.3835 | -104.3312
9.3873 | -104.3342
9.391 | -104.3366
9.3947 | -104.3364
9.3984 | -104.3335
9.4021 | -104.3281
9.4059 | -104.3218
9.4096 | -104.3144
9.4133 | -104.3029
9.417 | -104.2854
9.4207 | -104.2636
9.4245 | -104.2378
9.4282 | -104.2103
9.4319 | -104.1782
9.4356 | -104.1472
9.4393 | -104.1173
9.443 | -104.0865
9.4468 | -104.0507
9.4505 | -104.019
9.4542 | -103.9912
9.4579 | -103.9623
9.4616 | -103.9365
9.4654 | -103.9173
9.4691 | -103.8984
9.4728 | -103.8823
9.4765 | -103.8747
9.4802 | -103.875
9.484 | -103.8776
9.4877 | -103.882
9.4914 | -103.8901
9.4951 | -103.9074
9.4988 | -103.9268
9.5026 | -103.9525
9.5063 | -103.9818
9.51 | -104.0114
9.5137 | -104.0416
9.5174 | -104.0754
9.5212 | -104.1155
9.5249 | -104.1562
9.5286 | -104.2018
9.5323 | -104.2497
9.536 | -104.2929
9.5397 | -104.3345
9.5435 | -104.3739
9.5472 | -104.4114
9.5509 | -104.4465
9.5546 | -104.4756
9.5583 | -104.4967
9.5621 | -104.5148
9.5658 | -104.5264
9.5695 | -104.5323
9.5732 | -104.5344
9.5769 | -104.5314
9.5807 | -104.5288
9.5844 | -104.5281
9.5881 | -104.5271
9.5918 | -104.5236
9.5955 | -104.5279
9.5993 | -104.528
9.603 | -104.5248
9.6067 | -104.5296
9.6104 | -104.5367
9.6141 | -104.5431
9.6179 | -104.5509
9.6216 | -104.5639
9.6253 | -104.5783
9.629 | -104.5962
9.6327 | -104.6138
9.6364 | -104.6289
9.6402 | -104.6424
9.6439 | -104.6542
9.6476 | -104.668
9.6513 | -104.6815
9.655 | -104.6704
9.6588 | -104.6716
9.6625 | -104.6687
9.6662 | -104.6656
9.6699 | -104.6583
9.6736 | -104.654
9.6774 | -104.6501
9.6811 | -104.6448
9.6848 | -104.6373
9.6885 | -104.6288
9.6922 | -104.622
9.696 | -104.6231
9.6997 | -104.6244
9.7034 | -104.6208
9.7071 | -104.6231
9.7108 | -104.6329
9.7146 | -104.6413
9.7183 | -104.6487
9.722 | -104.6566
9.7257 | -104.6639
9.7294 | -104.671
9.7331 | -104.6834
9.7369 | -104.695
9.7406 | -104.7086
9.7443 | -104.7198
9.748 | -104.7276
9.7517 | -104.7356
9.7555 | -104.7439
9.7592 | -104.7553
9.7629 | -104.7671
9.7666 | -104.7805
9.7703 | -104.794
9.7741 | -104.8058
9.7778 | -104.8204
9.7815 | -104.8352
9.7852 | -104.8484
9.7889 | -104.8594
9.7927 | -104.8706
9.7964 | -104.884
9.8001 | -104.9027
9.8038 | -104.9139
9.8075 | -104.9217
9.8113 | -104.9282
9.815 | -104.933
9.8187 | -104.9324
9.8224 | -104.9326
9.8261 | -104.9293
9.8298 | -104.9271
9.8336 | -104.9259
9.8373 | -104.9237
9.841 | -104.9174
9.8447 | -104.9189
9.8484 | -104.9218
9.8522 | -104.9228
9.8559 | -104.9278
9.8596 | -104.9383
9.8633 | -104.951
9.867 | -104.965
9.8708 | -104.9808
9.8745 | -105.0017
9.8782 | -105.0199
9.8819 | -105.0388
9.8856 | -105.057
9.8894 | -105.0759
9.8931 | -105.0935
9.8968 | -105.1075
9.9005 | -105.1197
9.9042 | -105.1288
9.9079 | -105.1363
9.9117 | -105.1375
9.9154 | -105.1358
9.9191 | -105.1321
9.9228 | -105.122
9.9265 | -105.1129
9.9303 | -105.105
9.934 | -105.0895
9.9377 | -105.0692
9.9414 | -105.0494
9.9451 | -105.0236
9.9489 | -104.994
9.9526 | -104.9633
9.9563 | -104.9314
9.96 | -104.8936
9.9637 | -104.8569
9.9675 | -104.8191
9.9712 | -104.7836
9.9749 | -104.7436
9.9786 | -104.7053
9.9823 | -104.6697
9.9861 | -104.6369
9.9898 | -104.6046
9.9935 | -104.5797
9.9972 | -104.5593
10.0009 | -104.5395
10.0046 | -104.5238
10.0084 | -104.5187
10.0121 | -104.518
10.0158 | -104.5235
10.0195 | -104.5323
10.0232 | -104.5489
10.027 | -104.5687
10.0307 | -104.5942
10.0344 | -104.6212
10.0381 | -104.6477
10.0418 | -104.6749
10.0456 | -104.6983
10.0493 | -104.7196
10.053 | -104.7415
10.0567 | -104.7644
10.0604 | -104.7818
10.0642 | -104.796
10.0679 | -104.8071
10.0716 | -104.8111
10.0753 | -104.8092
10.079 | -104.8061
10.0828 | -104.8003
10.0865 | -104.7878
10.0902 | -104.7742
10.0939 | -104.7606
10.0976 | -104.7465
10.1013 | -104.7288
10.1051 | -104.7116
10.1088 | -104.6947
10.1125 | -104.6769
10.1162 | -104.6616
10.1199 | -104.6503
10.1237 | -104.6357
10.1274 | -104.6219
10.1311 | -104.6077
10.1348 | -104.5948
10.1385 | -104.5818
10.1423 | -104.5706
10.146 | -104.564
10.1497 | -104.5615
10.1534 | -104.5526
10.1571 | -104.5478
10.1609 | -104.5371
10.1646 | -104.53
10.1683 | -104.5249
10.172 | -104.5156
10.1757 | -104.5112
10.1795 | -104.5086
10.1832 | -104.5006
10.1869 | -104.4898
10.1906 | -104.4825
10.1943 | -104.4726
10.198 | -104.4649
10.2018 | -104.455
10.2055 | -104.4445
10.2092 | -104.435
10.2129 | -104.425
10.2166 | -104.4189
10.2204 | -104.4164
10.2241 | -104.4163
10.2278 | -104.4182
10.2315 | -104.4239
10.2352 | -104.435
10.239 | -104.4455
10.2427 | -104.4589
10.2464 | -104.475
10.2501 | -104.4947
10.2538 | -104.5131
10.2576 | -104.5289
10.2613 | -104.5459
10.265 | -104.5633
10.2687 | -104.5816
10.2724 | -104.6011
10.2762 | -104.6206
10.2799 | -104.6419
10.2836 | -104.6615
10.2873 | -104.679
10.291 | -104.6952
10.2947 | -104.7096
10.2985 | -104.7187
10.3022 | -104.73
10.3059 | -104.7376
10.3096 | -104.7468
10.3133 | -104.7496
10.3171 | -104.7504
10.3208 | -104.7561
10.3245 | -104.756
10.3282 | -104.7554
10.3319 | -104.7595
10.3357 | -104.7668
10.3394 | -104.769
10.3431 | -104.7751
10.3468 | -104.7792
10.3505 | -104.7825
10.3543 | -104.789
10.358 | -104.7978
10.3617 | -104.809
10.3654 | -104.8221
10.3691 | -104.8334
10.3728 | -104.8426
10.3766 | -104.8507
10.3803 | -104.8613
10.384 | -104.8731
10.3877 | -104.8827
10.3914 | -104.892
10.3952 | -104.9026
10.3989 | -104.9105
10.4026 | -104.9172
10.4063 | -104.9248
10.41 | -104.9305
10.4138 | -104.9327
10.4175 | -104.9316
10.4212 | -104.9284
10.4249 | -104.9219
10.4286 | -104.916
10.4324 | -104.9107
10.4361 | -104.9033
10.4398 | -104.899
10.4435 | -104.8982
10.4472 | -104.9012
10.451 | -104.9067
10.4547 | -104.9127
10.4584 | -104.9195
10.4621 | -104.9292
10.4658 | -104.9414
10.4695 | -104.9491
10.4733 | -104.9566
10.477 | -104.9658
10.4807 | -104.9757
10.4844 | -104.9827
10.4881 | -104.9854
10.4919 | -104.9896
10.4956 | -104.9895
10.4993 | -104.9903
10.503 | -104.9902
10.5067 | -104.9801
10.5105 | -104.9682
10.5142 | -104.9526
10.5179 | -104.9323
10.5216 | -104.9149
10.5253 | -104.8958
10.5291 | -104.8765
10.5328 | -104.8541
10.5365 | -104.829
10.5402 | -104.8077
10.5439 | -104.787
10.5477 | -104.766
10.5514 | -104.7481
10.5551 | -104.7321
10.5588 | -104.7157
10.5625 | -104.6991
10.5662 | -104.6906
10.57 | -104.6855
10.5737 | -104.6805
10.5774 | -104.6792
10.5811 | -104.6791
10.5848 | -104.6774
10.5886 | -104.6788
10.5923 | -104.6839
10.596 | -104.6906
10.5997 | -104.6962
10.6034 | -104.7076
10.6072 | -104.7245
10.6109 | -104.7439
10.6146 | -104.7652
10.6183 | -104.7844
10.622 | -104.8076
10.6258 | -104.8338
10.6295 | -104.8617
10.6332 | -104.8858
10.6369 | -104.9106
10.6406 | -104.9363
10.6444 | -104.9578
10.6481 | -104.9784
10.6518 | -104.9987
10.6555 | -105.0152
10.6592 | -105.0299
10.6629 | -105.0406
10.6667 | -105.0506
10.6704 | -105.0619
10.6741 | -105.0662
10.6778 | -105.0674
10.6815 | -105.0691
10.6853 | -105.0697
10.689 | -105.0685
10.6927 | -105.0648
10.6964 | -105.0614
10.7001 | -105.0543
10.7039 | -105.0499
10.7076 | -105.0486
10.7113 | -105.0514
10.715 | -105.0533
10.7187 | -105.0515
10.7225 | -105.0471
10.7262 | -105.0459
10.7299 | -105.0444
10.7336 | -105.0441
10.7373 | -105.0429
10.7411 | -105.0433
10.7448 | -105.0435
10.7485 | -105.0386
10.7522 | -105.0341
10.7559 | -105.0327
10.7596 | -105.027
10.7634 | -105.0195
10.7671 | -105.0134
10.7708 | -105.0054
10.7745 | -104.9985
10.7782 | -104.9901
10.782 | -104.9809
10.7857 | -104.9714
10.7894 | -104.9671
10.7931 | -104.9657
10.7968 | -104.9637
10.8006 | -104.9645
10.8043 | -104.9656
10.808 | -104.9665
10.8117 | -104.9723
10.8154 | -104.9806
10.8192 | -104.9888
10.8229 | -104.9948
10.8266 | -104.9988
10.8303 | -105.0052
10.834 | -105.0136
10.8377 | -105.0181
10.8415 | -105.022
10.8452 | -105.0293
10.8489 | -105.0374
10.8526 | -105.0434
10.8563 | -105.049
10.8601 | -105.0523
10.8638 | -105.0524
10.8675 | -105.0521
10.8712 | -105.0524
10.8749 | -105.0506
10.8787 | -105.0517
10.8824 | -105.052
10.8861 | -105.0478
10.8898 | -105.0465
10.8935 | -105.0446
10.8973 | -105.0433
10.901 | -105.0419
10.9047 | -105.0399
10.9084 | -105.0408
10.9121 | -105.0417
10.9159 | -105.0408
10.9196 | -105.0406
10.9233 | -105.0391
10.927 | -105.0389
10.9307 | -105.0356
10.9344 | -105.03
10.9382 | -105.0275
10.9419 | -105.0193
10.9456 | -105.0102
10.9493 | -104.9997
10.953 | -104.9863
10.9568 | -104.9695
10.9605 | -104.953
10.9642 | -104.9381
10.9679 | -104.9243
10.9716 | -104.9099
10.9754 | -104.8915
10.9791 | -104.8785
10.9828 | -104.8713
10.9865 | -104.8661
10.9902 | -104.8627
10.994 | -104.8649
10.9977 | -104.8696
11.0014 | -104.8731
11.0051 | -104.878
11.0088 | -104.8887
11.0126 | -104.9042
11.0163 | -104.9259
11.02 | -104.9493
11.0237 | -104.9745
11.0274 | -105.0025
11.0311 | -105.0332
11.0349 | -105.0649
11.0386 | -105.0995
11.0423 | -105.1336
11.046 | -105.1681
11.0497 | -105.197
11.0535 | -105.222
11.0572 | -105.2441
11.0609 | -105.2648
11.0646 | -105.284
11.0683 | -105.2954
11.0721 | -105.3037
11.0758 | -105.3092
11.0795 | -105.3126
11.0832 | -105.3132
11.0869 | -105.3069
11.0907 | -105.301
11.0944 | -105.2971
11.0981 | -105.2841
11.1018 | -105.2641
11.1055 | -105.2411
11.1093 | -105.2147
11.113 | -105.1929
11.1167 | -105.1749
11.1204 | -105.1543
11.1241 | -105.1325
11.1278 | -105.1116
11.1316 | -105.0921
11.1353 | -105.0784
11.139 | -105.0589
11.1427 | -105.0395
11.1464 | -105.0211
11.1502 | -105.006
11.1539 | -104.9917
11.1576 | -104.9807
11.1613 | -104.9722
11.165 | -104.9637
11.1688 | -104.9563
11.1725 | -104.9505
11.1762 | -104.9427
11.1799 | -104.9366
11.1836 | -104.9364
11.1874 | -104.9375
11.1911 | -104.9368
11.1948 | -104.9375
11.1985 | -104.9381
11.2022 | -104.9424
11.206 | -104.9512
11.2097 | -104.9625
11.2134 | -104.9707
11.2171 | -104.9806
11.2208 | -104.9866
11.2245 | -104.9899
11.2283 | -104.9985
11.232 | -105.0039
11.2357 | -105.0071
11.2394 | -105.0106
11.2431 | -105.0138
11.2469 | -105.0153
11.2506 | -105.0178
11.2543 | -105.0191
11.258 | -105.0231
11.2617 | -105.0296
11.2655 | -105.0323
11.2692 | -105.0401
11.2729 | -105.0471
11.2766 | -105.0466
11.2803 | -105.0453
11.2841 | -105.0462
11.2878 | -105.0481
11.2915 | -105.0516
11.2952 | -105.0564
11.2989 | -105.0638
11.3026 | -105.0736
11.3064 | -105.0836
11.3101 | -105.0958
11.3138 | -105.1061
11.3175 | -105.1177
11.3212 | -105.1306
11.325 | -105.1432
11.3287 | -105.154
11.3324 | -105.162
11.3361 | -105.1687
11.3398 | -105.1768
11.3436 | -105.188
11.3473 | -105.1934
11.351 | -105.1985
11.3547 | -105.208
11.3584 | -105.2171
11.3622 | -105.2194
11.3659 | -105.2201
11.3696 | -105.2237
11.3733 | -105.2287
11.377 | -105.2363
11.3808 | -105.2432
11.3845 | -105.2517
11.3882 | -105.2622
11.3919 | -105.2717
11.3956 | -105.2808
11.3993 | -105.2877
11.4031 | -105.2941
11.4068 | -105.3005
11.4105 | -105.3073
11.4142 | -105.3141
11.4179 | -105.3207
11.4217 | -105.3274
11.4254 | -105.3336
11.4291 | -105.3348
11.4328 | -105.3373
11.4365 | -105.34
11.4403 | -105.343
11.444 | -105.3492
11.4477 | -105.3541
11.4514 | -105.3556
11.4551 | -105.3554
11.4589 | -105.3542
11.4626 | -105.3529
11.4663 | -105.3465
11.47 | -105.3397
11.4737 | -105.3325
11.4775 | -105.3244
11.4812 | -105.3129
11.4849 | -105.3
11.4886 | -105.2908
11.4923 | -105.2833
11.496 | -105.2728
11.4998 | -105.2617
11.5035 | -105.2524
11.5072 | -105.2433
11.5109 | -105.2325
11.5146 | -105.2255
11.5184 | -105.2168
11.5221 | -105.2113
11.5258 | -105.2057
11.5295 | -105.2049
11.5332 | -105.2036
11.537 | -105.1965
11.5407 | -105.1878
11.5444 | -105.181
11.5481 | -105.1743
11.5518 | -105.1651
11.5556 | -105.1525
11.5593 | -105.1363
11.563 | -105.1207
11.5667 | -105.1097
11.5704 | -105.0975
11.5742 | -105.084
11.5779 | -105.0721
11.5816 | -105.0572
11.5853 | -105.0408
11.589 | -105.0305
11.5927 | -105.0234
11.5965 | -105.02
11.6002 | -105.0128
11.6039 | -105.0069
11.6076 | -105.002
11.6113 | -105.0002
11.6151 | -105.0019
11.6188 | -105.0063
11.6225 | -105.0149
11.6262 | -105.0245
11.6299 | -105.0361
11.6337 | -105.0516
11.6374 | -105.0686
11.6411 | -105.0898
11.6448 | -105.1116
11.6485 | -105.1337
11.6523 | -105.1579
11.656 | -105.1813
11.6597 | -105.2054
11.6634 | -105.232
11.6671 | -105.2628
11.6709 | -105.2892
11.6746 | -105.3109
11.6783 | -105.3324
11.682 | -105.3512
11.6857 | -105.3713
11.6894 | -105.3875
11.6932 | -105.4025
11.6969 | -105.4131
11.7006 | -105.4213
11.7043 | -105.4249
11.708 | -105.425
11.7118 | -105.4255
11.7155 | -105.4226
11.7192 | -105.4194
11.7229 | -105.4171
11.7266 | -105.4136
11.7304 | -105.4082
11.7341 | -105.3957
11.7378 | -105.3795
11.7415 | -105.3617
11.7452 | -105.3439
11.749 | -105.3291
11.7527 | -105.3156
11.7564 | -105.2969
11.7601 | -105.2819
11.7638 | -105.2692
11.7675 | -105.2571
11.7713 | -105.2437
11.775 | -105.2344
11.7787 | -105.2236
11.7824 | -105.211
11.7861 | -105.2019
11.7899 | -105.194
11.7936 | -105.1859
11.7973 | -105.181
11.801 | -105.1815
11.8047 | -105.1765
11.8085 | -105.1733
11.8122 | -105.1683
11.8159 | -105.167
11.8196 | -105.171
11.8233 | -105.1732
11.8271 | -105.1737
11.8308 | -105.1771
11.8345 | -105.1784
11.8382 | -105.1737
11.8419 | -105.176
11.8457 | -105.1783
11.8494 | -105.1752
11.8531 | -105.1751
11.8568 | -105.172
11.8605 | -105.167
11.8642 | -105.1634
11.868 | -105.1624
11.8717 | -105.1625
11.8754 | -105.1625
11.8791 | -105.1595
11.8828 | -105.1533
11.8866 | -105.1499
11.8903 | -105.146
11.894 | -105.1416
11.8977 | -105.1404
11.9014 | -105.1425
11.9052 | -105.1447
11.9089 | -105.1466
11.9126 | -105.1483
11.9163 | -105.149
11.92 | -105.1469
11.9238 | -105.1425
11.9275 | -105.1432
11.9312 | -105.1377
11.9349 | -105.1321
11.9386 | -105.1279
11.9424 | -105.1254
11.9461 | -105.1212
11.9498 | -105.1171
11.9535 | -105.1151
11.9572 | -105.1132
11.9609 | -105.1106
11.9647 | -105.1112
11.9684 | -105.1124
11.9721 | -105.1142
11.9758 | -105.116
11.9795 | -105.1215
11.9833 | -105.1255
11.987 | -105.1319
11.9907 | -105.1398
11.9944 | -105.1442
11.9981 | -105.1509
12.0019 | -105.162
12.0056 | -105.171
12.0093 | -105.1824
12.013 | -105.1979
12.0167 | -105.2093
12.0205 | -105.2206
12.0242 | -105.2364
12.0279 | -105.2491
12.0316 | -105.2582
12.0353 | -105.2641
12.0391 | -105.2686
12.0428 | -105.2753
12.0465 | -105.2801
12.0502 | -105.2871
12.0539 | -105.2931
12.0576 | -105.2945
12.0614 | -105.2936
12.0651 | -105.2926
12.0688 | -105.2895
12.0725 | -105.285
12.0762 | -105.2792
12.08 | -105.2757
12.0837 | -105.2731
12.0874 | -105.2719
12.0911 | -105.2676
12.0948 | -105.2649
12.0986 | -105.2631
12.1023 | -105.2586
12.106 | -105.2602
12.1097 | -105.2621
12.1134 | -105.2624
12.1172 | -105.2627
12.1209 | -105.267
12.1246 | -105.2711
12.1283 | -105.2777
12.132 | -105.2794
12.1358 | -105.2851
12.1395 | -105.2942
12.1432 | -105.3035
12.1469 | -105.3165
12.1506 | -105.3288
12.1543 | -105.3396
12.1581 | -105.3547
12.1618 | -105.3704
12.1655 | -105.3852
12.1692 | -105.4043
12.1729 | -105.4225
12.1767 | -105.4414
12.1804 | -105.4627
12.1841 | -105.4818
12.1878 | -105.4986
12.1915 | -105.517
12.1953 | -105.5334
12.199 | -105.5463
12.2027 | -105.5585
12.2064 | -105.5763
12.2101 | -105.5933
12.2139 | -105.6072
12.2176 | -105.6159
12.2213 | -105.6238
12.225 | -105.6318
12.2287 | -105.6386
12.2325 | -105.6431
12.2362 | -105.6436
12.2399 | -105.6447
12.2436 | -105.6446
12.2473 | -105.6412
12.251 | -105.6359
12.2548 | -105.6284
12.2585 | -105.6176
12.2622 | -105.6055
12.2659 | -105.5957
12.2696 | -105.583
12.2734 | -105.569
12.2771 | -105.5522
12.2808 | -105.5308
12.2845 | -105.5084
12.2882 | -105.4904
12.292 | -105.4715
12.2957 | -105.4507
12.2994 | -105.4323
12.3031 | -105.4105
12.3068 | -105.3941
12.3106 | -105.3751
12.3143 | -105.355
12.318 | -105.3342
12.3217 | -105.3096
12.3254 | -105.2871
12.3291 | -105.2662
12.3329 | -105.2437
12.3366 | -105.2218
12.3403 | -105.204
12.344 | -105.184
12.3477 | -105.1659
12.3515 | -105.1481
12.3552 | -105.1308
12.3589 | -105.1154
12.3626 | -105.1037
12.3663 | -105.093
12.3701 | -105.0816
12.3738 | -105.0747
12.3775 | -105.0723
12.3812 | -105.0685
12.3849 | -105.062
12.3887 | -105.0615
12.3924 | -105.0627
12.3961 | -105.0661
12.3998 | -105.0726
12.4035 | -105.0782
12.4073 | -105.0851
12.411 | -105.0915
12.4147 | -105.0985
12.4184 | -105.1046
12.4221 | -105.1104
12.4258 | -105.1171
12.4296 | -105.1251
12.4333 | -105.1329
12.437 | -105.1397
12.4407 | -105.1449
12.4444 | -105.1504
12.4482 | -105.1563
12.4519 | -105.1616
12.4556 | -105.1686
12.4593 | -105.1766
12.463 | -105.1849
12.4668 | -105.1895
12.4705 | -105.194
12.4742 | -105.2013
12.4779 | -105.2104
12.4816 | -105.223
12.4854 | -105.234
12.4891 | -105.2447
12.4928 | -105.2605
12.4965 | -105.2797
12.5002 | -105.3003
12.504 | -105.3212
12.5077 | -105.3363
12.5114 | -105.3545
12.5151 | -105.3729
12.5188 | -105.3931
12.5225 | -105.4162
12.5263 | -105.438
12.53 | -105.4596
12.5337 | -105.4798
12.5374 | -105.4971
12.5411 | -105.5123
12.5449 | -105.5229
12.5486 | -105.5239
12.5523 | -105.5257
12.556 | -105.5278
12.5597 | -105.5256
12.5635 | -105.5249
12.5672 | -105.522
12.5709 | -105.5132
12.5746 | -105.5014
12.5783 | -105.487
12.5821 | -105.4708
12.5858 | -105.4548
12.5895 | -105.4385
12.5932 | -105.4229
12.5969 | -105.4076
12.6007 | -105.3897
12.6044 | -105.3716
12.6081 | -105.355
12.6118 | -105.3353
12.6155 | -105.3167
12.6192 | -105.3053
12.623 | -105.2964
12.6267 | -105.2858
12.6304 | -105.2767
12.6341 | -105.2722
12.6378 | -105.2692
12.6416 | -105.2708
12.6453 | -105.2758
12.649 | -105.2827
12.6527 | -105.2896
12.6564 | -105.2956
12.6602 | -105.3025
12.6639 | -105.3082
12.6676 | -105.3167
12.6713 | -105.3259
12.675 | -105.3318
12.6788 | -105.338
12.6825 | -105.3473
12.6862 | -105.3554
12.6899 | -105.3673
12.6936 | -105.3761
12.6974 | -105.3829
12.7011 | -105.3886
12.7048 | -105.3904
12.7085 | -105.3959
12.7122 | -105.4044
12.7159 | -105.4123
12.7197 | -105.4226
12.7234 | -105.4348
12.7271 | -105.4468
12.7308 | -105.4615
12.7345 | -105.4729
12.7383 | -105.4804
12.742 | -105.4876
12.7457 | -105.4939
12.7494 | -105.5004
12.7531 | -105.5071
12.7569 | -105.5175
12.7606 | -105.5257
12.7643 | -105.5284
12.768 | -105.5311
12.7717 | -105.5347
12.7755 | -105.5359
12.7792 | -105.5376
12.7829 | -105.5431
12.7866 | -105.5459
12.7903 | -105.5459
12.794 | -105.5526
12.7978 | -105.5597
12.8015 | -105.5611
12.8052 | -105.5606
12.8089 | -105.5669
12.8126 | -105.5729
12.8164 | -105.5773
12.8201 | -105.5826
12.8238 | -105.5857
12.8275 | -105.5844
12.8312 | -105.5837
12.835 | -105.5829
12.8387 | -105.5827
12.8424 | -105.5809
12.8461 | -105.5772
12.8498 | -105.5695
12.8536 | -105.5597
12.8573 | -105.5489
12.861 | -105.5377
12.8647 | -105.524
12.8684 | -105.5108
12.8722 | -105.4981
12.8759 | -105.4832
12.8796 | -105.4676
12.8833 | -105.4524
12.887 | -105.4402
12.8907 | -105.4317
12.8945 | -105.4216
12.8982 | -105.4069
12.9019 | -105.3958
12.9056 | -105.3835
12.9093 | -105.3708
12.9131 | -105.3628
12.9168 | -105.3555
12.9205 | -105.3452
12.9242 | -105.3354
12.9279 | -105.3296
12.9317 | -105.3262
12.9354 | -105.3228
12.9391 | -105.3186
12.9428 | -105.3117
12.9465 | -105.3039
12.9503 | -105.3006
12.954 | -105.2969
12.9577 | -105.2912
12.9614 | -105.2854
12.9651 | -105.2753
12.9689 | -105.2659
12.9726 | -105.2615
12.9763 | -105.2584
12.98 | -105.2515
12.9837 | -105.2441
12.9874 | -105.2401
12.9912 | -105.239
12.9949 | -105.2387
12.9986 | -105.2406
13.0023 | -105.2445
13.006 | -105.2486
13.0098 | -105.2548
13.0135 | -105.2654
13.0172 | -105.2771
13.0209 | -105.2921
13.0246 | -105.3033
13.0284 | -105.3158
13.0321 | -105.3304
13.0358 | -105.346
13.0395 | -105.3623
13.0432 | -105.3777
13.047 | -105.3921
13.0507 | -105.4086
13.0544 | -105.4239
13.0581 | -105.4372
13.0618 | -105.4524
13.0656 | -105.4718
13.0693 | -105.49
13.073 | -105.5075
13.0767 | -105.5219
13.0804 | -105.535
13.0841 | -105.5492
13.0879 | -105.5658
13.0916 | -105.5822
13.0953 | -105.5994
13.099 | -105.612
13.1027 | -105.6244
13.1065 | -105.639
13.1102 | -105.6509
13.1139 | -105.6604
13.1176 | -105.6667
13.1213 | -105.6725
13.1251 | -105.6796
13.1288 | -105.6879
13.1325 | -105.6982
13.1362 | -105.7106
13.1399 | -105.7178
13.1437 | -105.721
13.1474 | -105.7242
13.1511 | -105.7263
13.1548 | -105.7288
13.1585 | -105.7251
13.1623 | -105.7184
13.166 | -105.711
13.1697 | -105.7016
13.1734 | -105.6894
13.1771 | -105.6781
13.1808 | -105.6617
13.1846 | -105.6408
13.1883 | -105.6199
13.192 | -105.5958
13.1957 | -105.5662
13.1994 | -105.5342
13.2032 | -105.5004
13.2069 | -105.4649
13.2106 | -105.4262
13.2143 | -105.3872
13.218 | -105.3498
13.2218 | -105.3109
13.2255 | -105.2689
13.2292 | -105.222
13.2329 | -105.1762
13.2366 | -105.1342
13.2404 | -105.0939
13.2441 | -105.0545
13.2478 | -105.0126
13.2515 | -104.9708
13.2552 | -104.9282
13.2589 | -104.8892
13.2627 | -104.8527
13.2664 | -104.8175
13.2701 | -104.7835
13.2738 | -104.7493
13.2775 | -104.7185
13.2813 | -104.6918
13.285 | -104.6679
13.2887 | -104.6459
13.2924 | -104.6252
13.2961 | -104.6073
13.2999 | -104.5944
13.3036 | -104.5847
13.3073 | -104.5754
13.311 | -104.5696
13.3147 | -104.5693
13.3185 | -104.5705
13.3222 | -104.5737
13.3259 | -104.5812
13.3296 | -104.595
13.3333 | -104.6126
13.3371 | -104.6316
13.3408 | -104.6525
13.3445 | -104.6761
13.3482 | -104.7051
13.3519 | -104.734
13.3556 | -104.7627
13.3594 | -104.7946
13.3631 | -104.8301
13.3668 | -104.8596
13.3705 | -104.8913
13.3742 | -104.9245
13.378 | -104.9561
13.3817 | -104.9886
13.3854 | -105.0211
13.3891 | -105.0527
13.3928 | -105.0807
13.3966 | -105.1062
13.4003 | -105.1292
13.404 | -105.1509
13.4077 | -105.1668
13.4114 | -105.1684
13.4152 | -105.1801
13.4189 | -105.1903
13.4226 | -105.2021
13.4263 | -105.2105
13.43 | -105.2138
13.4338 | -105.2131
13.4375 | -105.2127
13.4412 | -105.2142
13.4449 | -105.2176
13.4486 | -105.2219
13.4523 | -105.2231
13.4561 | -105.2224
13.4598 | -105.2261
13.4635 | -105.2305
13.4672 | -105.234
13.4709 | -105.2364
13.4747 | -105.2376
13.4784 | -105.2384
13.4821 | -105.2432
13.4858 | -105.2496
13.4895 | -105.2563
13.4933 | -105.2631
13.497 | -105.2714
13.5007 | -105.2816
13.5044 | -105.2866
13.5081 | -105.2891
13.5119 | -105.295
13.5156 | -105.3013
13.5193 | -105.3062
13.523 | -105.3079
13.5267 | -105.3089
13.5305 | -105.3068
13.5342 | -105.3023
13.5379 | -105.2977
13.5416 | -105.2929
13.5453 | -105.2899
13.549 | -105.2885
13.5528 | -105.2846
13.5565 | -105.2789
13.5602 | -105.2741
13.5639 | -105.2693
13.5676 | -105.2632
13.5714 | -105.2592
13.5751 | -105.2535
13.5788 | -105.2505
13.5825 | -105.2482
13.5862 | -105.2461
13.59 | -105.2448
13.5937 | -105.2438
13.5974 | -105.2441
13.6011 | -105.2451
13.6048 | -105.2473
13.6086 | -105.2471
13.6123 | -105.248
13.616 | -105.2504
13.6197 | -105.251
13.6234 | -105.2533
13.6272 | -105.2554
13.6309 | -105.2601
13.6346 | -105.2674
13.6383 | -105.2747
13.642 | -105.2799
13.6457 | -105.2851
13.6495 | -105.2938
13.6532 | -105.3064
13.6569 | -105.3151
13.6606 | -105.3239
13.6643 | -105.337
13.6681 | -105.3482
13.6718 | -105.3592
13.6755 | -105.3702
13.6792 | -105.3815
13.6829 | -105.3929
13.6867 | -105.4033
13.6904 | -105.4118
13.6941 | -105.4191
13.6978 | -105.4274
13.7015 | -105.4396
13.7053 | -105.4455
13.709 | -105.4492
13.7127 | -105.4515
13.7164 | -105.4555
13.7201 | -105.4597
13.7238 | -105.461
13.7276 | -105.4669
13.7313 | -105.469
13.735 | -105.4679
13.7387 | -105.4645
13.7424 | -105.4568
13.7462 | -105.4444
13.7499 | -105.4356
13.7536 | -105.4267
13.7573 | -105.4182
13.761 | -105.4057
13.7648 | -105.3925
13.7685 | -105.3791
13.7722 | -105.3639
13.7759 | -105.3471
13.7796 | -105.3285
13.7834 | -105.3122
13.7871 | -105.2966
13.7908 | -105.2824
13.7945 | -105.2682
13.7982 | -105.2535
13.802 | -105.2405
13.8057 | -105.2316
13.8094 | -105.2228
13.8131 | -105.2171
13.8168 | -105.213
13.8205 | -105.2086
13.8243 | -105.2099
13.828 | -105.2143
13.8317 | -105.2176
13.8354 | -105.2244
13.8391 | -105.2372
13.8429 | -105.2516
13.8466 | -105.2669
13.8503 | -105.2807
13.854 | -105.2992
13.8577 | -105.3202
13.8615 | -105.3409
13.8652 | -105.3657
13.8689 | -105.3893
13.8726 | -105.4117
13.8763 | -105.4328
13.8801 | -105.4517
13.8838 | -105.4714
13.8875 | -105.4947
13.8912 | -105.5162
13.8949 | -105.5361
13.8987 | -105.5553
13.9024 | -105.5765
13.9061 | -105.5916
13.9098 | -105.6028
13.9135 | -105.6099
13.9172 | -105.6157
13.921 | -105.6233
13.9247 | -105.6253
13.9284 | -105.6293
13.9321 | -105.6324
13.9358 | -105.6328
13.9396 | -105.6335
13.9433 | -105.6343
13.947 | -105.6308
13.9507 | -105.6261
13.9544 | -105.6193
13.9582 | -105.612
13.9619 | -105.609
13.9656 | -105.6071
13.9693 | -105.6042
13.973 | -105.6032
13.9768 | -105.5996
13.9805 | -105.5963
13.9842 | -105.5962
13.9879 | -105.5935
13.9916 | -105.5924
13.9954 | -105.5921
13.9991 | -105.5934
14.0028 | -105.5943
14.0065 | -105.5941
14.0102 | -105.5922
14.0139 | -105.5868
14.0177 | -105.5832
14.0214 | -105.582
14.0251 | -105.5779
14.0288 | -105.5714
14.0325 | -105.5659
14.0363 | -105.5604
14.04 | -105.5543
14.0437 | -105.5505
14.0474 | -105.549
14.0511 | -105.5521
14.0549 | -105.5551
14.0586 | -105.5599
14.0623 | -105.5622
14.066 | -105.5635
14.0697 | -105.5706
14.0735 | -105.5849
14.0772 | -105.6001
14.0809 | -105.6159
14.0846 | -105.631
14.0883 | -105.6481
14.0921 | -105.6647
14.0958 | -105.6813
14.0995 | -105.6985
14.1032 | -105.7182
14.1069 | -105.7398
14.1106 | -105.7625
14.1144 | -105.7816
14.1181 | -105.7985
14.1218 | -105.8122
14.1255 | -105.8237
14.1292 | -105.8366
14.133 | -105.8504
14.1367 | -105.8632
14.1404 | -105.876
14.1441 | -105.8876
14.1478 | -105.8956
14.1516 | -105.9026
14.1553 | -105.9078
14.159 | -105.9127
14.1627 | -105.916
14.1664 | -105.916
14.1702 | -105.9123
14.1739 | -105.9064
14.1776 | -105.9016
14.1813 | -105.8978
14.185 | -105.8898
14.1887 | -105.8714
14.1925 | -105.8538
14.1962 | -105.8374
14.1999 | -105.8184
14.2036 | -105.7976
14.2073 | -105.7781
14.2111 | -105.7583
14.2148 | -105.7388
14.2185 | -105.72
14.2222 | -105.6954
14.2259 | -105.6703
14.2297 | -105.6472
14.2334 | -105.625
14.2371 | -105.6045
14.2408 | -105.5853
14.2445 | -105.5639
14.2483 | -105.5451
14.252 | -105.5294
14.2557 | -105.5128
14.2594 | -105.4986
14.2631 | -105.488
14.2669 | -105.4821
14.2706 | -105.4763
14.2743 | -105.4693
14.278 | -105.4611
14.2817 | -105.4561
14.2854 | -105.4562
14.2892 | -105.4594
14.2929 | -105.4613
14.2966 | -105.4672
14.3003 | -105.4735
14.304 | -105.4818
14.3078 | -105.4914
14.3115 | -105.5024
14.3152 | -105.5152
14.3189 | -105.5302
14.3226 | -105.5465
14.3264 | -105.56
14.3301 | -105.5749
14.3338 | -105.5844
14.3375 | -105.5914
14.3412 | -105.5988
14.345 | -105.6082
14.3487 | -105.6162
14.3524 | -105.6232
14.3561 | -105.627
14.3598 | -105.6306
14.3636 | -105.6334
14.3673 | -105.6363
14.371 | -105.635
14.3747 | -105.6326
14.3784 | -105.6322
14.3821 | -105.6324
14.3859 | -105.6322
14.3896 | -105.6321
14.3933 | -105.6288
14.397 | -105.6253
14.4007 | -105.6202
14.4045 | -105.6173
14.4082 | -105.6134
14.4119 | -105.6078
14.4156 | -105.601
14.4193 | -105.5932
14.4231 | -105.5864
14.4268 | -105.5839
14.4305 | -105.5781
14.4342 | -105.5706
14.4379 | -105.5643
14.4417 | -105.5554
14.4454 | -105.5435
14.4491 | -105.5293
14.4528 | -105.5129
14.4565 | -105.4953
14.4603 | -105.4783
14.464 | -105.4596
14.4677 | -105.4398
14.4714 | -105.4213
14.4751 | -105.4046
14.4788 | -105.3899
14.4826 | -105.3754
14.4863 | -105.3657
14.49 | -105.352
14.4937 | -105.3385
14.4974 | -105.3297
14.5012 | -105.3257
14.5049 | -105.3195
14.5086 | -105.3137
14.5123 | -105.3107
14.516 | -105.3094
14.5198 | -105.3111
14.5235 | -105.3136
14.5272 | -105.3178
14.5309 | -105.3232
14.5346 | -105.3337
14.5384 | -105.3442
14.5421 | -105.3557
14.5458 | -105.364
14.5495 | -105.3723
14.5532 | -105.3793
14.557 | -105.386
14.5607 | -105.3931
14.5644 | -105.4006
14.5681 | -105.4079
14.5718 | -105.4101
14.5755 | -105.4088
14.5793 | -105.4091
14.583 | -105.4079
14.5867 | -105.4068
14.5904 | -105.4032
14.5941 | -105.4013
14.5979 | -105.4003
14.6016 | -105.4019
14.6053 | -105.3991
14.609 | -105.3934
14.6127 | -105.3857
14.6165 | -105.3809
14.6202 | -105.375
14.6239 | -105.364
14.6276 | -105.3543
14.6313 | -105.3458
14.6351 | -105.3362
14.6388 | -105.3264
14.6425 | -105.3164
14.6462 | -105.3043
14.6499 | -105.2937
14.6536 | -105.2846
14.6574 | -105.2762
14.6611 | -105.2659
14.6648 | -105.2534
14.6685 | -105.2408
14.6722 | -105.2301
14.676 | -105.2198
14.6797 | -105.2099
14.6834 | -105.1992
14.6871 | -105.1868
14.6908 | -105.1736
14.6946 | -105.1598
14.6983 | -105.1495
14.702 | -105.1392
14.7057 | -105.1285
14.7094 | -105.1177
14.7132 | -105.105
14.7169 | -105.099
14.7206 | -105.0957
14.7243 | -105.0894
14.728 | -105.0823
14.7318 | -105.0761
14.7355 | -105.0722
14.7392 | -105.0696
14.7429 | -105.0673
14.7466 | -105.0649
14.7503 | -105.0629
14.7541 | -105.0614
14.7578 | -105.0619
14.7615 | -105.0657
14.7652 | -105.0695
14.7689 | -105.0762
14.7727 | -105.0861
14.7764 | -105.0955
14.7801 | -105.1056
14.7838 | -105.1209
14.7875 | -105.1399
14.7913 | -105.1626
14.795 | -105.1885
14.7987 | -105.2139
14.8024 | -105.2431
14.8061 | -105.2746
14.8099 | -105.3085
14.8136 | -105.3469
14.8173 | -105.3884
14.821 | -105.43
14.8247 | -105.4738
14.8285 | -105.515
14.8322 | -105.5587
14.8359 | -105.6088
14.8396 | -105.656
14.8433 | -105.6984
14.847 | -105.7385
14.8508 | -105.7777
14.8545 | -105.8166
14.8582 | -105.8533
14.8619 | -105.8858
14.8656 | -105.9132
14.8694 | -105.9395
14.8731 | -105.9632
14.8768 | -105.9809
14.8805 | -105.996
14.8842 | -106.0058
14.888 | -106.0126
14.8917 | -106.0182
14.8954 | -106.0205
14.8991 | -106.0184
14.9028 | -106.0165
14.9066 | -106.0102
14.9103 | -105.9971
14.914 | -105.9849
14.9177 | -105.9667
14.9214 | -105.944
14.9252 | -105.9226
14.9289 | -105.8958
14.9326 | -105.8667
14.9363 | -105.8387
14.94 | -105.8108
14.9437 | -105.7817
14.9475 | -105.7514
14.9512 | -105.7215
14.9549 | -105.6911
14.9586 | -105.659
14.9623 | -105.6276
14.9661 | -105.5981
14.9698 | -105.5676
14.9735 | -105.5362
14.9772 | -105.5068
14.9809 | -105.4768
14.9847 | -105.4446
14.9884 | -105.4187
14.9921 | -105.3938
14.9958 | -105.3695
14.9995 | -105.3461
15.0033 | -105.322
15.007 | -105.3015
15.0107 | -105.2849
15.0144 | -105.2713
15.0181 | -105.2595
15.0219 | -105.2618
15.0256 | -105.2547
15.0293 | -105.2511
15.033 | -105.2483
15.0367 | -105.2475
15.0404 | -105.2493
15.0442 | -105.2506
15.0479 | -105.2538
15.0516 | -105.2592
15.0553 | -105.2666
15.059 | -105.275
15.0628 | -105.2848
15.0665 | -105.2961
15.0702 | -105.3059
15.0739 | -105.318
15.0776 | -105.3318
15.0814 | -105.3472
15.0851 | -105.3609
15.0888 | -105.3747
15.0925 | -105.393
15.0962 | -105.4096
15.1 | -105.4262
15.1037 | -105.441
15.1074 | -105.454
15.1111 | -105.4694
15.1148 | -105.484
15.1186 | -105.4965
15.1223 | -105.5129
15.126 | -105.5303
15.1297 | -105.5448
15.1334 | -105.5598
15.1371 | -105.5736
15.1409 | -105.5853
15.1446 | -105.5968
15.1483 | -105.6082
15.152 | -105.6197
15.1557 | -105.6309
15.1595 | -105.6401
15.1632 | -105.6479
15.1669 | -105.6546
15.1706 | -105.6627
15.1743 | -105.6709
15.1781 | -105.6777
15.1818 | -105.6852
15.1855 | -105.6935
15.1892 | -105.6992
15.1929 | -105.7016
15.1967 | -105.7038
15.2004 | -105.7076
15.2041 | -105.7096
15.2078 | -105.7107
15.2115 | -105.7113
15.2152 | -105.7137
15.219 | -105.7137
15.2227 | -105.7133
15.2264 | -105.7146
15.2301 | -105.7134
15.2338 | -105.7127
15.2376 | -105.7132
15.2413 | -105.7156
15.245 | -105.7182
15.2487 | -105.7185
15.2524 | -105.7196
15.2562 | -105.7207
15.2599 | -105.7207
15.2636 | -105.7227
15.2673 | -105.7239
15.271 | -105.7244
15.2748 | -105.7256
15.2785 | -105.7262
15.2822 | -105.7242
15.2859 | -105.7227
15.2896 | -105.7222
15.2934 | -105.7217
15.2971 | -105.723
15.3008 | -105.7239
15.3045 | -105.7229
15.3082 | -105.7216
15.3119 | -105.7175
15.3157 | -105.7142
15.3194 | -105.7096
15.3231 | -105.7073
15.3268 | -105.7062
15.3305 | -105.7025
15.3343 | -105.701
15.338 | -105.7007
15.3417 | -105.698
15.3454 | -105.6944
15.3491 | -105.6898
15.3529 | -105.6865
15.3566 | -105.6812
15.3603 | -105.6766
15.364 | -105.6717
15.3677 | -105.6661
15.3715 | -105.6574
15.3752 | -105.6486
15.3789 | -105.6437
15.3826 | -105.6368
15.3863 | -105.63
15.3901 | -105.6206
15.3938 | -105.6108
15.3975 | -105.6026
15.4012 | -105.5923
15.4049 | -105.5831
15.4086 | -105.5747
15.4124 | -105.569
15.4161 | -105.5617
15.4198 | -105.5539
15.4235 | -105.5452
15.4272 | -105.5367
15.431 | -105.5273
15.4347 | -105.5186
15.4384 | -105.5129
15.4421 | -105.5062
15.4458 | -105.4972
15.4496 | -105.4902
15.4533 | -105.484
15.457 | -105.4773
15.4607 | -105.4687
15.4644 | -105.4607
15.4682 | -105.4548
15.4719 | -105.4534
15.4756 | -105.4518
15.4793 | -105.448
15.483 | -105.4451
15.4868 | -105.4404
15.4905 | -105.4339
15.4942 | -105.4267
15.4979 | -105.4221
15.5016 | -105.4223
15.5053 | -105.4181
15.5091 | -105.4105
15.5128 | -105.4021
15.5165 | -105.3943
15.5202 | -105.3873
15.5239 | -105.3844
15.5277 | -105.3792
15.5314 | -105.3745
15.5351 | -105.3693
15.5388 | -105.3643
15.5425 | -105.3578
15.5463 | -105.3495
15.55 | -105.3432
15.5537 | -105.3385
15.5574 | -105.3369
15.5611 | -105.3297
15.5649 | -105.3229
15.5686 | -105.3179
15.5723 | -105.3153
15.576 | -105.3151
15.5797 | -105.3182
15.5834 | -105.3208
15.5872 | -105.3214
15.5909 | -105.3227
15.5946 | -105.3265
15.5983 | -105.3267
15.602 | -105.3254
15.6058 | -105.3294
15.6095 | -105.3347
15.6132 | -105.3425
15.6169 | -105.352
15.6206 | -105.359
15.6244 | -105.3662
15.6281 | -105.3742
15.6318 | -105.3829
15.6355 | -105.3916
15.6392 | -105.4037
15.643 | -105.4168
15.6467 | -105.427
15.6504 | -105.4368
15.6541 | -105.4482
15.6578 | -105.4592
15.6616 | -105.4692
15.6653 | -105.481
15.669 | -105.4946
15.6727 | -105.5086
15.6764 | -105.5224
15.6801 | -105.5369
15.6839 | -105.5523
15.6876 | -105.5667
15.6913 | -105.5795
15.695 | -105.5904
15.6987 | -105.6027
15.7025 | -105.6175
15.7062 | -105.6304
15.7099 | -105.64
15.7136 | -105.6503
15.7173 | -105.6628
15.7211 | -105.674
15.7248 | -105.6831
15.7285 | -105.6944
15.7322 | -105.7087
15.7359 | -105.7188
15.7397 | -105.7269
15.7434 | -105.7353
15.7471 | -105.7413
15.7508 | -105.7456
15.7545 | -105.7504
15.7583 | -105.755
15.762 | -105.7598
15.7657 | -105.7639
15.7694 | -105.7684
15.7731 | -105.7721
15.7768 | -105.7736
15.7806 | -105.7759
15.7843 | -105.7776
15.788 | -105.7798
15.7917 | -105.7837
15.7954 | -105.7845
15.7992 | -105.7856
15.8029 | -105.7867
15.8066 | -105.7868
15.8103 | -105.7881
15.814 | -105.7918
15.8178 | -105.7917
15.8215 | -105.7921
15.8252 | -105.7974
15.8289 | -105.7988
15.8326 | -105.799
15.8364 | -105.798
15.8401 | -105.7978
15.8438 | -105.7946
15.8475 | -105.7926
15.8512 | -105.7919
15.855 | -105.7901
15.8587 | -105.7868
15.8624 | -105.7832
15.8661 | -105.7798
15.8698 | -105.7745
15.8735 | -105.7692
15.8773 | -105.767
15.881 | -105.7636
15.8847 | -105.7565
15.8884 | -105.748
15.8921 | -105.7386
15.8959 | -105.7282
15.8996 | -105.7195
15.9033 | -105.712
15.907 | -105.7019
15.9107 | -105.693
15.9145 | -105.6835
15.9182 | -105.674
15.9219 | -105.663
15.9256 | -105.6515
15.9293 | -105.6388
15.9331 | -105.6294
15.9368 | -105.622
15.9405 | -105.6124
15.9442 | -105.5992
15.9479 | -105.586
15.9517 | -105.5731
15.9554 | -105.5606
15.9591 | -105.5474
15.9628 | -105.5367
15.9665 | -105.5242
15.9702 | -105.51
15.974 | -105.4973
15.9777 | -105.4864
15.9814 | -105.4747
15.9851 | -105.4626
15.9888 | -105.4512
15.9926 | -105.4411
15.9963 | -105.4282
16 | -105.4139
16.0037 | -105.3984
16.0074 | -105.3855
16.0112 | -105.3736
16.0149 | -105.3616
16.0186 | -105.3476
16.0223 | -105.3374
16.026 | -105.3253
16.0298 | -105.3151
16.0335 | -105.3081
16.0372 | -105.298
16.0409 | -105.2863
16.0446 | -105.2797
16.0483 | -105.2741
16.0521 | -105.2652
16.0558 | -105.2578
16.0595 | -105.2511
16.0632 | -105.2442
16.0669 | -105.2349
16.0707 | -105.2266
16.0744 | -105.22
16.0781 | -105.2128
16.0818 | -105.206
16.0855 | -105.1991
16.0893 | -105.1952
16.093 | -105.1894
16.0967 | -105.179
16.1004 | -105.1698
16.1041 | -105.1624
16.1079 | -105.1568
16.1116 | -105.1504
16.1153 | -105.144
16.119 | -105.1389
16.1227 | -105.1296
16.1265 | -105.1212
16.1302 | -105.1146
16.1339 | -105.1072
16.1376 | -105.1009
16.1413 | -105.0937
16.145 | -105.0892
16.1488 | -105.0869
16.1525 | -105.0841
16.1562 | -105.0818
16.1599 | -105.0793
16.1636 | -105.0743
16.1674 | -105.071
16.1711 | -105.0716
16.1748 | -105.0731
16.1785 | -105.0732
16.1822 | -105.0725
16.186 | -105.0728
16.1897 | -105.0736
16.1934 | -105.0762
16.1971 | -105.0798
16.2008 | -105.0863
16.2046 | -105.0936
16.2083 | -105.1047
16.212 | -105.1157
16.2157 | -105.1264
16.2194 | -105.1385
16.2232 | -105.1493
16.2269 | -105.1623
16.2306 | -105.1785
16.2343 | -105.195
16.238 | -105.209
16.2417 | -105.2225
16.2455 | -105.2359
16.2492 | -105.2499
16.2529 | -105.2633
16.2566 | -105.2787
16.2603 | -105.2906
16.2641 | -105.301
16.2678 | -105.3151
16.2715 | -105.3316
16.2752 | -105.3428
16.2789 | -105.3522
16.2827 | -105.3642
16.2864 | -105.3782
16.2901 | -105.3894
16.2938 | -105.4016
16.2975 | -105.4137
16.3013 | -105.4231
16.305 | -105.4304
16.3087 | -105.4383
16.3124 | -105.4486
16.3161 | -105.4591
16.3199 | -105.4665
16.3236 | -105.4765
16.3273 | -105.4887
16.331 | -105.5029
16.3347 | -105.5158
16.3384 | -105.5274
16.3422 | -105.535
16.3459 | -105.5423
16.3496 | -105.5512
16.3533 | -105.5603
16.357 | -105.5736
16.3608 | -105.5864
16.3645 | -105.5947
16.3682 | -105.6036
16.3719 | -105.6134
16.3756 | -105.6225
16.3794 | -105.6312
16.3831 | -105.6376
16.3868 | -105.644
16.3905 | -105.6507
16.3942 | -105.6568
16.398 | -105.664
16.4017 | -105.6677
16.4054 | -105.6713
16.4091 | -105.6756
16.4128 | -105.6795
16.4166 | -105.6828
16.4203 | -105.6864
16.424 | -105.6888
16.4277 | -105.6905
16.4314 | -105.6912
16.4351 | -105.6863
16.4389 | -105.6795
16.4426 | -105.6724
16.4463 | -105.6653
16.45 | -105.6595
16.4537 | -105.6542
16.4575 | -105.6513
16.4612 | -105.6496
16.4649 | -105.6459
16.4686 | -105.6419
16.4723 | -105.6391
16.4761 | -105.6356
16.4798 | -105.6336
16.4835 | -105.632
16.4872 | -105.6283
16.4909 | -105.6234
16.4947 | -105.6186
16.4984 | -105.6136
16.5021 | -105.6092
16.5058 | -105.6056
16.5095 | -105.6055
16.5132 | -105.6033
16.517 | -105.5996
16.5207 | -105.5944
16.5244 | -105.5902
16.5281 | -105.5842
16.5318 | -105.5803
16.5356 | -105.5785
16.5393 | -105.5767
16.543 | -105.5725
16.5467 | -105.5656
16.5504 | -105.5595
16.5542 | -105.5529
16.5579 | -105.5459
16.5616 | -105.5418
16.5653 | -105.5376
16.569 | -105.5317
16.5728 | -105.5279
16.5765 | -105.5234
16.5802 | -105.5203
16.5839 | -105.5166
16.5876 | -105.5114
16.5914 | -105.5058
16.5951 | -105.4984
16.5988 | -105.4948
16.6025 | -105.4919
16.6062 | -105.4891
16.6099 | -105.4894
16.6137 | -105.4875
16.6174 | -105.4867
16.6211 | -105.4894
16.6248 | -105.4915
16.6285 | -105.4989
16.6323 | -105.5137
16.636 | -105.5216
16.6397 | -105.5303
16.6434 | -105.5378
16.6471 | -105.5442
16.6509 | -105.553
16.6546 | -105.5653
16.6583 | -105.5762
16.662 | -105.5879
16.6657 | -105.5982
16.6695 | -105.6095
16.6732 | -105.6245
16.6769 | -105.6382
16.6806 | -105.6491
16.6843 | -105.6604
16.6881 | -105.6705
16.6918 | -105.6791
16.6955 | -105.685
16.6992 | -105.6905
16.7029 | -105.6946
16.7066 | -105.697
16.7104 | -105.6981
16.7141 | -105.6981
16.7178 | -105.698
16.7215 | -105.6956
16.7252 | -105.6922
16.729 | -105.6894
16.7327 | -105.6874
16.7364 | -105.6833
16.7401 | -105.675
16.7438 | -105.6619
16.7476 | -105.6494
16.7513 | -105.6391
16.755 | -105.6303
16.7587 | -105.619
16.7624 | -105.6028
16.7662 | -105.5866
16.7699 | -105.5728
16.7736 | -105.5588
16.7773 | -105.5437
16.781 | -105.5291
16.7848 | -105.5147
16.7885 | -105.5011
16.7922 | -105.4881
16.7959 | -105.4737
16.7996 | -105.4601
16.8033 | -105.4481
16.8071 | -105.4371
16.8108 | -105.4246
16.8145 | -105.4133
16.8182 | -105.4022
16.8219 | -105.3929
16.8257 | -105.3851
16.8294 | -105.3778
16.8331 | -105.3707
16.8368 | -105.3618
16.8405 | -105.3548
16.8443 | -105.3491
16.848 | -105.342
16.8517 | -105.3383
16.8554 | -105.3365
16.8591 | -105.3318
16.8629 | -105.3308
16.8666 | -105.3277
16.8703 | -105.3247
16.874 | -105.3259
16.8777 | -105.3278
16.8815 | -105.3315
16.8852 | -105.3358
16.8889 | -105.3433
16.8926 | -105.3517
16.8963 | -105.3595
16.9 | -105.3692
16.9038 | -105.3786
16.9075 | -105.3902
16.9112 | -105.4029
16.9149 | -105.4147
16.9186 | -105.4256
16.9224 | -105.4363
16.9261 | -105.4426
16.9298 | -105.4508
16.9335 | -105.4611
16.9372 | -105.4727
16.941 | -105.4816
16.9447 | -105.4906
16.9484 | -105.5023
16.9521 | -105.5133
16.9558 | -105.5214
16.9596 | -105.5267
16.9633 | -105.5319
16.967 | -105.538
16.9707 | -105.5441
16.9744 | -105.5484
16.9781 | -105.5526
16.9819 | -105.558
16.9856 | -105.5648
16.9893 | -105.567
16.993 | -105.5672
16.9967 | -105.568
17.0005 | -105.5675
17.0042 | -105.568
17.0079 | -105.5686
17.0116 | -105.5629
17.0153 | -105.5573
17.0191 | -105.552
17.0228 | -105.5451
17.0265 | -105.5387
17.0302 | -105.5327
17.0339 | -105.5256
17.0377 | -105.5184
17.0414 | -105.5097
17.0451 | -105.5011
17.0488 | -105.4925
17.0525 | -105.4852
17.0563 | -105.4761
17.06 | -105.4655
17.0637 | -105.4551
17.0674 | -105.4451
17.0711 | -105.4351
17.0748 | -105.4261
17.0786 | -105.4172
17.0823 | -105.4095
17.086 | -105.402
17.0897 | -105.3931
17.0934 | -105.3845
17.0972 | -105.376
17.1009 | -105.3683
17.1046 | -105.3608
17.1083 | -105.3556
17.112 | -105.3532
17.1158 | -105.3471
17.1195 | -105.3422
17.1232 | -105.338
17.1269 | -105.3341
17.1306 | -105.3279
17.1344 | -105.3211
17.1381 | -105.3158
17.1418 | -105.312
17.1455 | -105.3093
17.1492 | -105.3047
17.153 | -105.2994
17.1567 | -105.2954
17.1604 | -105.295
17.1641 | -105.2936
17.1678 | -105.2873
17.1715 | -105.2862
17.1753 | -105.2799
17.179 | -105.2723
17.1827 | -105.2658
17.1864 | -105.2609
17.1901 | -105.2563
17.1939 | -105.2504
17.1976 | -105.2437
17.2013 | -105.239
17.205 | -105.2356
17.2087 | -105.2306
17.2125 | -105.2249
17.2162 | -105.2186
17.2199 | -105.2121
17.2236 | -105.2071
17.2273 | -105.2038
17.2311 | -105.2005
17.2348 | -105.1991
17.2385 | -105.1978
17.2422 | -105.1953
17.2459 | -105.1956
17.2497 | -105.1988
17.2534 | -105.2005
17.2571 | -105.2034
17.2608 | -105.2047
17.2645 | -105.2055
17.2682 | -105.2097
17.272 | -105.2128
17.2757 | -105.2152
17.2794 | -105.2204
17.2831 | -105.2249
17.2868 | -105.2323
17.2906 | -105.2399
17.2943 | -105.2466
17.298 | -105.2533
17.3017 | -105.2626
17.3054 | -105.2729
17.3092 | -105.2822
17.3129 | -105.2899
17.3166 | -105.2969
17.3203 | -105.3047
17.324 | -105.3106
17.3278 | -105.3133
17.3315 | -105.3171
17.3352 | -105.3227
17.3389 | -105.3276
17.3426 | -105.3313
17.3464 | -105.3364
17.3501 | -105.3384
17.3538 | -105.338
17.3575 | -105.3352
17.3612 | -105.3338
17.3649 | -105.3326
17.3687 | -105.3299
17.3724 | -105.3251
17.3761 | -105.3243
17.3798 | -105.3257
17.3835 | -105.3256
17.3873 | -105.3251
17.391 | -105.3256
17.3947 | -105.3284
17.3984 | -105.3303
17.4021 | -105.3306
17.4059 | -105.3312
17.4096 | -105.3334
17.4133 | -105.3376
17.417 | -105.343
17.4207 | -105.3502
17.4245 | -105.3599
17.4282 | -105.3688
17.4319 | -105.3758
17.4356 | -105.3809
17.4393 | -105.3881
17.443 | -105.3986
17.4468 | -105.4096
17.4505 | -105.4209
17.4542 | -105.4309
17.4579 | -105.442
17.4616 | -105.4537
17.4654 | -105.4644
17.4691 | -105.4726
17.4728 | -105.4796
17.4765 | -105.484
17.4802 | -105.4915
17.484 | -105.4975
17.4877 | -105.5031
17.4914 | -105.5067
17.4951 | -105.5089
17.4988 | -105.5075
17.5026 | -105.5059
17.5063 | -105.4988
17.51 | -105.4915
17.5137 | -105.4836
17.5174 | -105.4713
17.5212 | -105.4576
17.5249 | -105.4415
17.5286 | -105.4247
17.5323 | -105.4078
17.536 | -105.3875
17.5397 | -105.3648
17.5435 | -105.345
17.5472 | -105.3227
17.5509 | -105.2972
17.5546 | -105.2689
17.5583 | -105.2432
17.5621 | -105.2182
17.5658 | -105.1912
17.5695 | -105.1634
17.5732 | -105.1394
17.5769 | -105.1172
17.5807 | -105.0929
17.5844 | -105.0707
17.5881 | -105.0478
17.5918 | -105.0247
17.5955 | -105.0029
17.5993 | -104.9854
17.603 | -104.9702
17.6067 | -104.9574
17.6104 | -104.9424
17.6141 | -104.9272
17.6179 | -104.9178
17.6216 | -104.9126
17.6253 | -104.9046
17.629 | -104.8973
17.6327 | -104.8938
17.6364 | -104.8923
17.6402 | -104.8912
17.6439 | -104.8916
17.6476 | -104.8945
17.6513 | -104.9016
17.655 | -104.9087
17.6588 | -104.9167
17.6625 | -104.9246
17.6662 | -104.9343
17.6699 | -104.9462
17.6736 | -104.9578
17.6774 | -104.9684
17.6811 | -104.9818
17.6848 | -104.9966
17.6885 | -105.0091
17.6922 | -105.0185
17.696 | -105.0298
17.6997 | -105.0414
17.7034 | -105.0472
17.7071 | -105.0562
17.7108 | -105.0627
17.7146 | -105.0664
17.7183 | -105.0704
17.722 | -105.0759
17.7257 | -105.0795
17.7294 | -105.0813
17.7331 | -105.0811
17.7369 | -105.0822
17.7406 | -105.0815
17.7443 | -105.0798
17.748 | -105.0783
17.7517 | -105.0752
17.7555 | -105.0694
17.7592 | -105.0692
17.7629 | -105.0655
17.7666 | -105.0587
17.7703 | -105.0519
17.7741 | -105.0446
17.7778 | -105.0376
17.7815 | -105.0262
17.7852 | -105.0151
17.7889 | -105.0046
17.7927 | -104.9937
17.7964 | -104.9819
17.8001 | -104.9701
17.8038 | -104.9585
17.8075 | -104.9493
17.8113 | -104.9388
17.815 | -104.9286
17.8187 | -104.9175
17.8224 | -104.9076
17.8261 | -104.8946
17.8298 | -104.8796
17.8336 | -104.869
17.8373 | -104.8579
17.841 | -104.8473
17.8447 | -104.8388
17.8484 | -104.8325
17.8522 | -104.8248
17.8559 | -104.8172
17.8596 | -104.8094
17.8633 | -104.8042
17.867 | -104.7986
17.8708 | -104.7894
17.8745 | -104.7812
17.8782 | -104.7746
17.8819 | -104.7683
17.8856 | -104.7609
17.8894 | -104.7516
17.8931 | -104.7438
17.8968 | -104.7373
17.9005 | -104.7325
17.9042 | -104.726
17.9079 | -104.7194
17.9117 | -104.7134
17.9154 | -104.7104
17.9191 | -104.7061
17.9228 | -104.7012
17.9265 | -104.6973
17.9303 | -104.6922
17.934 | -104.6882
17.9377 | -104.6866
17.9414 | -104.6848
17.9451 | -104.6844
17.9489 | -104.6849
17.9526 | -104.6852
17.9563 | -104.6876
17.96 | -104.6918
17.9637 | -104.6954
17.9675 | -104.7012
17.9712 | -104.7091
17.9749 | -104.7157
17.9786 | -104.7224
17.9823 | -104.729
17.9861 | -104.7376
17.9898 | -104.7449
17.9935 | -104.7514
17.9972 | -104.7605
18.0009 | -104.7694
18.0046 | -104.7799
18.0084 | -104.7926
18.0121 | -104.8046
18.0158 | -104.8131
18.0195 | -104.8226
18.0232 | -104.8329
18.027 | -104.8428
18.0307 | -104.8548
18.0344 | -104.8675
18.0381 | -104.8815
18.0418 | -104.8943
18.0456 | -104.9063
18.0493 | -104.9189
18.053 | -104.9316
18.0567 | -104.9447
18.0604 | -104.9575
18.0642 | -104.9691
18.0679 | -104.9784
18.0716 | -104.9858
18.0753 | -104.995
18.079 | -105.0042
18.0828 | -105.0141
18.0865 | -105.0237
18.0902 | -105.0326
18.0939 | -105.0417
18.0976 | -105.0519
18.1013 | -105.0602
18.1051 | -105.0652
18.1088 | -105.0709
18.1125 | -105.0785
18.1162 | -105.0863
18.1199 | -105.092
18.1237 | -105.0998
18.1274 | -105.1067
18.1311 | -105.1127
18.1348 | -105.1206
18.1385 | -105.1267
18.1423 | -105.1311
18.146 | -105.1365
18.1497 | -105.1441
18.1534 | -105.1501
18.1571 | -105.1553
18.1609 | -105.1628
18.1646 | -105.171
18.1683 | -105.1788
18.172 | -105.1857
18.1757 | -105.1941
18.1795 | -105.2012
18.1832 | -105.2107
18.1869 | -105.2229
18.1906 | -105.2336
18.1943 | -105.2417
18.198 | -105.2516
18.2018 | -105.26
18.2055 | -105.2646
18.2092 | -105.2702
18.2129 | -105.2785
18.2166 | -105.286
18.2204 | -105.2939
18.2241 | -105.2954
18.2278 | -105.2972
18.2315 | -105.3011
18.2352 | -105.3034
18.239 | -105.303
18.2427 | -105.288
18.2464 | -105.2835
18.2501 | -105.2783
18.2538 | -105.2704
18.2576 | -105.2594
18.2613 | -105.2482
18.265 | -105.2305
18.2687 | -105.2134
18.2724 | -105.1965
18.2762 | -105.1796
18.2799 | -105.1623
18.2836 | -105.1437
18.2873 | -105.122
18.291 | -105.0993
18.2947 | -105.0739
18.2985 | -105.0489
18.3022 | -105.0248
18.3059 | -104.9999
18.3096 | -104.975
18.3133 | -104.9499
18.3171 | -104.9244
18.3208 | -104.9003
18.3245 | -104.8758
18.3282 | -104.8548
18.3319 | -104.8341
18.3357 | -104.8146
18.3394 | -104.7968
18.3431 | -104.7793
18.3468 | -104.7638
18.3505 | -104.7454
18.3543 | -104.7275
18.358 | -104.712
18.3617 | -104.6991
18.3654 | -104.6893
18.3691 | -104.6794
18.3728 | -104.6702
18.3766 | -104.6638
18.3803 | -104.6585
18.384 | -104.6535
18.3877 | -104.6519
18.3914 | -104.6492
18.3952 | -104.646
18.3989 | -104.647
18.4026 | -104.6517
18.4063 | -104.6587
18.41 | -104.6617
18.4138 | -104.6635
18.4175 | -104.6678
18.4212 | -104.6719
18.4249 | -104.6761
18.4286 | -104.6808
18.4324 | -104.6866
18.4361 | -104.6921
18.4398 | -104.7004
18.4435 | -104.7073
18.4472 | -104.7121
18.451 | -104.7157
18.4547 | -104.7226
18.4584 | -104.7305
18.4621 | -104.7387
18.4658 | -104.7477
18.4695 | -104.7532
18.4733 | -104.7555
18.477 | -104.7574
18.4807 | -104.7612
18.4844 | -104.7671
18.4881 | -104.7705
18.4919 | -104.7714
18.4956 | -104.7729
18.4993 | -104.7749
18.503 | -104.7774
18.5067 | -104.7792
18.5105 | -104.7796
18.5142 | -104.7764
18.5179 | -104.7765
18.5216 | -104.7757
18.5253 | -104.776
18.5291 | -104.7803
18.5328 | -104.7808
18.5365 | -104.7783
18.5402 | -104.7734
18.5439 | -104.7681
18.5477 | -104.7605
18.5514 | -104.7529
18.5551 | -104.7437
18.5588 | -104.738
18.5625 | -104.7323
18.5662 | -104.724
18.57 | -104.7156
18.5737 | -104.7073
18.5774 | -104.696
18.5811 | -104.6845
18.5848 | -104.6742
18.5886 | -104.6648
18.5923 | -104.6517
18.596 | -104.6385
18.5997 | -104.6263
18.6034 | -104.6103
18.6072 | -104.5947
18.6109 | -104.5778
18.6146 | -104.5627
18.6183 | -104.5473
18.622 | -104.5318
18.6258 | -104.5171
18.6295 | -104.4994
18.6332 | -104.4823
18.6369 | -104.4643
18.6406 | -104.4469
18.6444 | -104.43
18.6481 | -104.414
18.6518 | -104.3983
18.6555 | -104.3854
18.6592 | -104.371
18.6629 | -104.3541
18.6667 | -104.3376
18.6704 | -104.3224
18.6741 | -104.3064
18.6778 | -104.2934
18.6815 | -104.2815
18.6853 | -104.2694
18.689 | -104.2589
18.6927 | -104.2505
18.6964 | -104.2404
18.7001 | -104.233
18.7039 | -104.2254
18.7076 | -104.2182
18.7113 | -104.2124
18.715 | -104.2084
18.7187 | -104.2015
18.7225 | -104.1915
18.7262 | -104.1848
18.7299 | -104.1792
18.7336 | -104.1742
18.7373 | -104.168
18.7411 | -104.159
18.7448 | -104.1499
18.7485 | -104.1426
18.7522 | -104.1346
18.7559 | -104.1266
18.7596 | -104.1183
18.7634 | -104.1102
18.7671 | -104.1025
18.7708 | -104.0943
18.7745 | -104.0844
18.7782 | -104.0752
18.782 | -104.0665
18.7857 | -104.0564
18.7894 | -104.0465
18.7931 | -104.0349
18.7968 | -104.0238
18.8006 | -104.0131
18.8043 | -104.0032
18.808 | -103.9936
18.8117 | -103.9817
18.8154 | -103.9706
18.8192 | -103.9598
18.8229 | -103.9487
18.8266 | -103.935
18.8303 | -103.919
18.834 | -103.906
18.8377 | -103.8929
18.8415 | -103.8779
18.8452 | -103.8648
18.8489 | -103.8497
18.8526 | -103.832
18.8563 | -103.816
18.8601 | -103.8009
18.8638 | -103.787
18.8675 | -103.7715
18.8712 | -103.7565
18.8749 | -103.7407
18.8787 | -103.7261
18.8824 | -103.7121
18.8861 | -103.7001
18.8898 | -103.6869
18.8935 | -103.6708
18.8973 | -103.6554
18.901 | -103.6416
18.9047 | -103.6287
18.9084 | -103.6153
18.9121 | -103.6035
18.9159 | -103.5909
18.9196 | -103.5792
18.9233 | -103.568
18.927 | -103.5589
18.9307 | -103.5527
18.9344 | -103.5466
18.9382 | -103.5393
18.9419 | -103.5361
18.9456 | -103.5319
18.9493 | -103.5245
18.953 | -103.5207
18.9568 | -103.5148
18.9605 | -103.5097
18.9642 | -103.5071
18.9679 | -103.5045
18.9716 | -103.5023
18.9754 | -103.4988
18.9791 | -103.4949
18.9828 | -103.4916
18.9865 | -103.4895
18.9902 | -103.4899
18.994 | -103.488
18.9977 | -103.4866
19.0014 | -103.488
19.0051 | -103.4893
19.0088 | -103.4915
19.0126 | -103.4933
19.0163 | -103.4914
19.02 | -103.4896
19.0237 | -103.4912
19.0274 | -103.4919
19.0311 | -103.4906
19.0349 | -103.4886
19.0386 | -103.4867
19.0423 | -103.486
19.046 | -103.4859
19.0497 | -103.4857
19.0535 | -103.4862
19.0572 | -103.4857
19.0609 | -103.4844
19.0646 | -103.4836
19.0683 | -103.4817
19.0721 | -103.4822
19.0758 | -103.4831
19.0795 | -103.4842
19.0832 | -103.4862
19.0869 | -103.4882
19.0907 | -103.4869
19.0944 | -103.4868
19.0981 | -103.4885
19.1018 | -103.491
19.1055 | -103.4918
19.1093 | -103.4938
19.113 | -103.496
19.1167 | -103.4982
19.1204 | -103.5017
19.1241 | -103.5054
19.1278 | -103.5079
19.1316 | -103.5114
19.1353 | -103.5144
19.139 | -103.5181
19.1427 | -103.5226
19.1464 | -103.526
19.1502 | -103.5307
19.1539 | -103.5359
19.1576 | -103.5401
19.1613 | -103.5443
19.165 | -103.5505
19.1688 | -103.5576
19.1725 | -103.5637
19.1762 | -103.5701
19.1799 | -103.5758
19.1836 | -103.5824
19.1874 | -103.5891
19.1911 | -103.5959
19.1948 | -103.6037
19.1985 | -103.6114
19.2022 | -103.6183
19.206 | -103.6244
19.2097 | -103.6306
19.2134 | -103.6373
19.2171 | -103.6461
19.2208 | -103.6582
19.2245 | -103.6669
19.2283 | -103.6781
19.232 | -103.6887
19.2357 | -103.6956
19.2394 | -103.7011
19.2431 | -103.7075
19.2469 | -103.7153
19.2506 | -103.7241
19.2543 | -103.7329
19.258 | -103.74
19.2617 | -103.7466
19.2655 | -103.7531
19.2692 | -103.7592
19.2729 | -103.7638
19.2766 | -103.7698
19.2803 | -103.7763
19.2841 | -103.7799
19.2878 | -103.7856
19.2915 | -103.7903
19.2952 | -103.7911
19.2989 | -103.7915
19.3026 | -103.7955
19.3064 | -103.7994
19.3101 | -103.8025
19.3138 | -103.804
19.3175 | -103.8071
19.3212 | -103.8072
19.325 | -103.8079
19.3287 | -103.8095
19.3324 | -103.8109
19.3361 | -103.814
19.3398 | -103.8172
19.3436 | -103.8204
19.3473 | -103.8208
19.351 | -103.8232
19.3547 | -103.8275
19.3584 | -103.834
19.3622 | -103.8396
19.3659 | -103.8457
19.3696 | -103.8522
19.3733 | -103.86
19.377 | -103.8696
19.3808 | -103.8803
19.3845 | -103.8899
19.3882 | -103.8995
19.3919 | -103.9093
19.3956 | -103.9196
19.3993 | -103.9309
19.4031 | -103.942
19.4068 | -103.9526
19.4105 | -103.9644
19.4142 | -103.9762
19.4179 | -103.9872
19.4217 | -104.0004
19.4254 | -104.0117
19.4291 | -104.0203
19.4328 | -104.028
19.4365 | -104.0364
19.4403 | -104.0469
19.444 | -104.0545
19.4477 | -104.0599
19.4514 | -104.0638
19.4551 | -104.0665
19.4589 | -104.0646
19.4626 | -104.064
19.4663 | -104.0627
19.47 | -104.0595
19.4737 | -104.0577
19.4775 | -104.0556
19.4812 | -104.0537
19.4849 | -104.0501
19.4886 | -104.0436
19.4923 | -104.0357
19.496 | -104.0259
19.4998 | -104.0186
19.5035 | -104.0098
19.5072 | -103.9991
19.5109 | -103.9919
19.5146 | -103.9838
19.5184 | -103.9733
19.5221 | -103.966
19.5258 | -103.9584
19.5295 | -103.9492
19.5332 | -103.9412
19.537 | -103.929
19.5407 | -103.9176
19.5444 | -103.9083
19.5481 | -103.9007
19.5518 | -103.8914
19.5556 | -103.8835
19.5593 | -103.8792
19.563 | -103.8767
19.5667 | -103.8742
19.5704 | -103.8699
19.5742 | -103.8697
19.5779 | -103.8692
19.5816 | -103.8707
19.5853 | -103.8742
19.589 | -103.8767
19.5927 | -103.8786
19.5965 | -103.8838
19.6002 | -103.8919
19.6039 | -103.901
19.6076 | -103.9097
19.6113 | -103.9194
19.6151 | -103.9289
19.6188 | -103.9379
19.6225 | -103.9482
19.6262 | -103.9588
19.6299 | -103.9668
19.6337 | -103.9762
19.6374 | -103.985
19.6411 | -103.9941
19.6448 | -104.005
19.6485 | -104.0134
19.6523 | -104.0202
19.656 | -104.0251
19.6597 | -104.0312
19.6634 | -104.0374
19.6671 | -104.0436
19.6709 | -104.0493
19.6746 | -104.0529
19.6783 | -104.055
19.682 | -104.0552
19.6857 | -104.0569
19.6894 | -104.057
19.6932 | -104.0567
19.6969 | -104.0561
19.7006 | -104.0544
19.7043 | -104.0514
19.708 | -104.0491
19.7118 | -104.0466
19.7155 | -104.0419
19.7192 | -104.0365
19.7229 | -104.0321
19.7266 | -104.0269
19.7304 | -104.0208
19.7341 | -104.0157
19.7378 | -104.0132
19.7415 | -104.0068
19.7452 | -103.9993
19.749 | -103.9923
19.7527 | -103.9841
19.7564 | -103.9764
19.7601 | -103.9707
19.7638 | -103.9638
19.7675 | -103.9577
19.7713 | -103.9513
19.775 | -103.9455
19.7787 | -103.9383
19.7824 | -103.9324
19.7861 | -103.9283
19.7899 | -103.9243
19.7936 | -103.9221
19.7973 | -103.9198
19.801 | -103.9151
19.8047 | -103.91
19.8085 | -103.9061
19.8122 | -103.9047
19.8159 | -103.9046
19.8196 | -103.9048
19.8233 | -103.905
19.8271 | -103.9047
19.8308 | -103.9041
19.8345 | -103.9029
19.8382 | -103.8998
19.8419 | -103.8971
19.8457 | -103.8974
19.8494 | -103.9023
19.8531 | -103.906
19.8568 | -103.9085
19.8605 | -103.911
19.8642 | -103.9113
19.868 | -103.9126
19.8717 | -103.9179
19.8754 | -103.9233
19.8791 | -103.9261
19.8828 | -103.9287
19.8866 | -103.9322
19.8903 | -103.9382
19.894 | -103.943
19.8977 | -103.9471
19.9014 | -103.9526
19.9052 | -103.9578
19.9089 | -103.9623
19.9126 | -103.9681
19.9163 | -103.9714
19.92 | -103.9752
19.9238 | -103.979
19.9275 | -103.9845
19.9312 | -103.9899
19.9349 | -103.9939
19.9386 | -103.9964
19.9424 | -103.9984
19.9461 | -104.0016
19.9498 | -104.0046
19.9535 | -104.0066
19.9572 | -104.0108
19.9609 | -104.014
19.9647 | -104.016
19.9684 | -104.0159
19.9721 | -104.0157
19.9758 | -104.0156
19.9795 | -104.0166
19.9833 | -104.0165
19.987 | -104.0168
19.9907 | -104.0172
19.9944 | -104.0181
19.9981 | -104.0189
20.0019 | -104.0196
20.0056 | -104.0204
20.0093 | -104.0206
20.013 | -104.0196
20.0167 | -104.0185
20.0205 | -104.0179
20.0242 | -104.0192
20.0279 | -104.0188
20.0316 | -104.0199
20.0353 | -104.022
20.0391 | -104.023
20.0428 | -104.0238
20.0465 | -104.0261
20.0502 | -104.0277
20.0539 | -104.0294
20.0576 | -104.0313
20.0614 | -104.0338
20.0651 | -104.0377
20.0688 | -104.0414
20.0725 | -104.0456
20.0762 | -104.0491
20.08 | -104.0547
20.0837 | -104.0616
20.0874 | -104.0696
20.0911 | -104.0757
20.0948 | -104.0796
20.0986 | -104.0863
20.1023 | -104.0936
20.106 | -104.0993
20.1097 | -104.1043
20.1134 | -104.1126
20.1172 | -104.1199
20.1209 | -104.1284
20.1246 | -104.1371
20.1283 | -104.1433
20.132 | -104.1468
20.1358 | -104.1498
20.1395 | -104.1538
20.1432 | -104.1569
20.1469 | -104.1589
20.1506 | -104.1614
20.1543 | -104.1636
20.1581 | -104.164
20.1618 | -104.1635
20.1655 | -104.1632
20.1692 | -104.1607
20.1729 | -104.1573
20.1767 | -104.1525
20.1804 | -104.1487
20.1841 | -104.142
20.1878 | -104.1324
20.1915 | -104.1231
20.1953 | -104.117
20.199 | -104.1103
20.2027 | -104.1006
20.2064 | -104.0908
20.2101 | -104.0819
20.2139 | -104.0706
20.2176 | -104.058
20.2213 | -104.0478
20.225 | -104.0389
20.2287 | -104.0281
20.2325 | -104.0176
20.2362 | -104.0084
20.2399 | -103.9974
20.2436 | -103.9828
20.2473 | -103.9699
20.251 | -103.9592
20.2548 | -103.9501
20.2585 | -103.9401
20.2622 | -103.9294
20.2659 | -103.918
20.2696 | -103.9081
20.2734 | -103.8994
20.2771 | -103.8898
20.2808 | -103.8798
20.2845 | -103.8713
20.2882 | -103.8607
20.292 | -103.8487
20.2957 | -103.8411
20.2994 | -103.8355
20.3031 | -103.8295
20.3068 | -103.8233
20.3106 | -103.8149
20.3143 | -103.8091
20.318 | -103.8019
20.3217 | -103.796
20.3254 | -103.7906
20.3291 | -103.7862
20.3329 | -103.783
20.3366 | -103.7792
20.3403 | -103.7764
20.344 | -103.7737
20.3477 | -103.7728
20.3515 | -103.7712
20.3552 | -103.7712
20.3589 | -103.7719
20.3626 | -103.7733
20.3663 | -103.7733
20.3701 | -103.7734
20.3738 | -103.7757
20.3775 | -103.7795
20.3812 | -103.7847
20.3849 | -103.7877
20.3887 | -103.7911
20.3924 | -103.7957
20.3961 | -103.8015
20.3998 | -103.8083
20.4035 | -103.8136
20.4073 | -103.8183
20.411 | -103.8229
20.4147 | -103.8268
20.4184 | -103.8312
20.4221 | -103.8365
20.4258 | -103.8417
20.4296 | -103.8473
20.4333 | -103.8532
20.437 | -103.8587
20.4407 | -103.8651
20.4444 | -103.8697
20.4482 | -103.8758
20.4519 | -103.8812
20.4556 | -103.8862
20.4593 | -103.8914
20.463 | -103.8958
20.4668 | -103.8994
20.4705 | -103.902
20.4742 | -103.9057
20.4779 | -103.9088
20.4816 | -103.9076
20.4854 | -103.9042
20.4891 | -103.8998
20.4928 | -103.8978
20.4965 | -103.895
20.5002 | -103.8899
20.504 | -103.8837
20.5077 | -103.8767
20.5114 | -103.8656
20.5151 | -103.8556
20.5188 | -103.8437
20.5225 | -103.8279
20.5263 | -103.8129
20.53 | -103.7988
20.5337 | -103.7835
20.5374 | -103.7679
20.5411 | -103.7508
20.5449 | -103.7329
20.5486 | -103.7147
20.5523 | -103.6964
20.556 | -103.6766
20.5597 | -103.6579
20.5635 | -103.6349
20.5672 | -103.6112
20.5709 | -103.5886
20.5746 | -103.5678
20.5783 | -103.5474
20.5821 | -103.5304
20.5858 | -103.5094
20.5895 | -103.4876
20.5932 | -103.4648
20.5969 | -103.4402
20.6007 | -103.4186
20.6044 | -103.3996
20.6081 | -103.3829
20.6118 | -103.3634
20.6155 | -103.3454
20.6192 | -103.3286
20.623 | -103.3129
20.6267 | -103.2981
20.6304 | -103.283
20.6341 | -103.2675
20.6378 | -103.2507
20.6416 | -103.2371
20.6453 | -103.2229
20.649 | -103.2087
20.6527 | -103.1962
20.6564 | -103.1882
20.6602 | -103.1798
20.6639 | -103.169
20.6676 | -103.1586
20.6713 | -103.1512
20.675 | -103.145
20.6788 | -103.1391
20.6825 | -103.1323
20.6862 | -103.1259
20.6899 | -103.1181
20.6936 | -103.1115
20.6974 | -103.1074
20.7011 | -103.1049
20.7048 | -103.1008
20.7085 | -103.0959
20.7122 | -103.0938
20.7159 | -103.0933
20.7197 | -103.091
20.7234 | -103.0887
20.7271 | -103.0849
20.7308 | -103.0828
20.7345 | -103.0818
20.7383 | -103.0831
20.742 | -103.0847
20.7457 | -103.0852
20.7494 | -103.0857
20.7531 | -103.0863
20.7569 | -103.0886
20.7606 | -103.0905
20.7643 | -103.0931
20.768 | -103.0969
20.7717 | -103.0991
20.7755 | -103.1042
20.7792 | -103.1069
20.7829 | -103.1099
20.7866 | -103.1125
20.7903 | -103.1132
20.794 | -103.1135
20.7978 | -103.1131
20.8015 | -103.1144
20.8052 | -103.1172
20.8089 | -103.1186
20.8126 | -103.1204
20.8164 | -103.1227
20.8201 | -103.1223
20.8238 | -103.1231
20.8275 | -103.1267
20.8312 | -103.1281
20.835 | -103.1289
20.8387 | -103.1282
20.8424 | -103.1272
20.8461 | -103.1268
20.8498 | -103.1289
20.8536 | -103.1292
20.8573 | -103.1326
20.861 | -103.1368
20.8647 | -103.1412
20.8684 | -103.1438
20.8722 | -103.1475
20.8759 | -103.1508
20.8796 | -103.1572
20.8833 | -103.1639
20.887 | -103.1734
20.8907 | -103.1844
20.8945 | -103.1939
20.8982 | -103.2011
20.9019 | -103.208
20.9056 | -103.2172
20.9093 | -103.2257
20.9131 | -103.2332
20.9168 | -103.2423
20.9205 | -103.2517
20.9242 | -103.2627
20.9279 | -103.2729
20.9317 | -103.284
20.9354 | -103.2929
20.9391 | -103.3015
20.9428 | -103.3119
20.9465 | -103.322
20.9503 | -103.3306
20.954 | -103.3371
20.9577 | -103.3447
20.9614 | -103.3519
20.9651 | -103.3588
20.9689 | -103.3643
20.9726 | -103.371
20.9763 | -103.377
20.98 | -103.3832
20.9837 | -103.3876
20.9874 | -103.3936
20.9912 | -103.3998
20.9949 | -103.4051
20.9986 | -103.4097
21.0023 | -103.4152
21.006 | -103.4194
21.0098 | -103.422
21.0135 | -103.4232
21.0172 | -103.4262
21.0209 | -103.4293
21.0246 | -103.434
21.0284 | -103.4366
21.0321 | -103.4377
21.0358 | -103.4373
21.0395 | -103.4376
21.0432 | -103.4397
21.047 | -103.4419
21.0507 | -103.4459
21.0544 | -103.4487
21.0581 | -103.4504
21.0618 | -103.4525
21.0656 | -103.4554
21.0693 | -103.4601
21.073 | -103.4636
21.0767 | -103.4706
21.0804 | -103.4765
21.0841 | -103.4809
21.0879 | -103.487
21.0916 | -103.4961
21.0953 | -103.5061
21.099 | -103.5143
21.1027 | -103.5229
21.1065 | -103.5295
21.1102 | -103.5375
21.1139 | -103.5479
21.1176 | -103.5586
21.1213 | -103.5669
21.1251 | -103.5741
21.1288 | -103.5826
21.1325 | -103.594
21.1362 | -103.6061
21.1399 | -103.619
21.1437 | -103.6289
21.1474 | -103.6362
21.1511 | -103.6467
21.1548 | -103.6577
21.1585 | -103.6692
21.1623 | -103.6795
21.166 | -103.6899
21.1697 | -103.7
21.1734 | -103.7081
21.1771 | -103.7136
21.1808 | -103.7193
21.1846 | -103.7269
21.1883 | -103.7348
21.192 | -103.7428
21.1957 | -103.7499
21.1994 | -103.7571
21.2032 | -103.7641
21.2069 | -103.7702
21.2106 | -103.7764
21.2143 | -103.7822
21.218 | -103.7883
21.2218 | -103.7935
21.2255 | -103.7983
21.2292 | -103.8016
21.2329 | -103.8059
21.2366 | -103.8134
21.2404 | -103.8176
21.2441 | -103.8208
21.2478 | -103.8253
21.2515 | -103.8289
21.2552 | -103.8324
21.2589 | -103.8357
21.2627 | -103.8403
21.2664 | -103.8457
21.2701 | -103.8498
21.2738 | -103.8522
21.2775 | -103.8562
21.2813 | -103.8578
21.285 | -103.8599
21.2887 | -103.8609
21.2924 | -103.8621
21.2961 | -103.8631
21.2999 | -103.8645
21.3036 | -103.8661
21.3073 | -103.8684
21.311 | -103.8694
21.3147 | -103.8724
21.3185 | -103.8757
21.3222 | -103.8766
21.3259 | -103.8769
21.3296 | -103.8786
21.3333 | -103.8807
21.3371 | -103.8828
21.3408 | -103.8836
21.3445 | -103.8858
21.3482 | -103.8876
21.3519 | -103.8884
21.3556 | -103.8892
21.3594 | -103.8923
21.3631 | -103.8935
21.3668 | -103.8929
21.3705 | -103.8944
21.3742 | -103.8964
21.378 | -103.8969
21.3817 | -103.8992
21.3854 | -103.9009
21.3891 | -103.9009
21.3928 | -103.9016
21.3966 | -103.9036
21.4003 | -103.9061
21.404 | -103.9076
21.4077 | -103.9104
21.4114 | -103.9108
21.4152 | -103.9108
21.4189 | -103.9117
21.4226 | -103.913
21.4263 | -103.9144
21.43 | -103.9165
21.4338 | -103.9165
21.4375 | -103.917
21.4412 | -103.9178
21.4449 | -103.9209
21.4486 | -103.9248
21.4523 | -103.9244
21.4561 | -103.9234
21.4598 | -103.9185
21.4635 | -103.9163
21.4672 | -103.915
21.4709 | -103.9136
21.4747 | -103.9126
21.4784 | -103.9111
21.4821 | -103.9115
21.4858 | -103.9124
21.4895 | -103.9098
21.4933 | -103.9061
21.497 | -103.9011
21.5007 | -103.8987
21.5044 | -103.8965
21.5081 | -103.8933
21.5119 | -103.8905
21.5156 | -103.8853
21.5193 | -103.8806
21.523 | -103.8758
21.5267 | -103.8699
21.5305 | -103.8637
21.5342 | -103.8569
21.5379 | -103.8483
21.5416 | -103.8397
21.5453 | -103.8337
21.549 | -103.8295
21.5528 | -103.8229
21.5565 | -103.8176
21.5602 | -103.8139
21.5639 | -103.8105
21.5676 | -103.8043
21.5714 | -103.7971
21.5751 | -103.7905
21.5788 | -103.7842
21.5825 | -103.7795
21.5862 | -103.7742
21.59 | -103.7705
21.5937 | -103.7676
21.5974 | -103.7624
21.6011 | -103.7574
21.6048 | -103.7513
21.6086 | -103.7471
21.6123 | -103.7437
21.616 | -103.7411
21.6197 | -103.7361
21.6234 | -103.7303
21.6272 | -103.7247
21.6309 | -103.7175
21.6346 | -103.7099
21.6383 | -103.7029
21.642 | -103.6937
21.6457 | -103.6842
21.6495 | -103.6752
21.6532 | -103.6674
21.6569 | -103.6594
21.6606 | -103.6512
21.6643 | -103.6424
21.6681 | -103.6334
21.6718 | -103.6253
21.6755 | -103.6161
21.6792 | -103.6053
21.6829 | -103.5946
21.6867 | -103.5822
21.6904 | -103.5713
21.6941 | -103.5628
21.6978 | -103.5526
21.7015 | -103.5402
21.7053 | -103.5265
21.709 | -103.5145
21.7127 | -103.501
21.7164 | -103.489
21.7201 | -103.4757
21.7238 | -103.4616
21.7276 | -103.4473
21.7313 | -103.4345
21.735 | -103.4201
21.7387 | -103.4063
21.7424 | -103.3942
21.7462 | -103.3796
21.7499 | -103.3646
21.7536 | -103.3505
21.7573 | -103.3377
21.761 | -103.3243
21.7648 | -103.312
21.7685 | -103.3022
21.7722 | -103.291
21.7759 | -103.276
21.7796 | -103.2601
21.7834 | -103.2464
21.7871 | -103.2341
21.7908 | -103.221
21.7945 | -103.2079
21.7982 | -103.1949
21.802 | -103.1846
21.8057 | -103.1745
21.8094 | -103.1637
21.8131 | -103.1531
21.8168 | -103.1421
21.8205 | -103.133
21.8243 | -103.1254
21.828 | -103.1163
21.8317 | -103.1075
21.8354 | -103.1002
21.8391 | -103.0951
21.8429 | -103.0895
21.8466 | -103.0827
21.8503 | -103.0749
21.854 | -103.0654
21.8577 | -103.0559
21.8615 | -103.0503
21.8652 | -103.0452
21.8689 | -103.0396
21.8726 | -103.0358
21.8763 | -103.0357
21.8801 | -103.0351
21.8838 | -103.0349
21.8875 | -103.0355
21.8912 | -103.0337
21.8949 | -103.0324
21.8987 | -103.0302
21.9024 | -103.029
21.9061 | -103.029
21.9098 | -103.0295
21.9135 | -103.0328
21.9172 | -103.0338
21.921 | -103.0363
21.9247 | -103.0389
21.9284 | -103.0416
21.9321 | -103.0456
21.9358 | -103.0523
21.9396 | -103.0609
21.9433 | -103.0686
21.947 | -103.0746
21.9507 | -103.0813
21.9544 | -103.0861
21.9582 | -103.0923
21.9619 | -103.0991
21.9656 | -103.106
21.9693 | -103.1119
21.973 | -103.1174
21.9768 | -103.124
21.9805 | -103.1289
21.9842 | -103.1334
21.9879 | -103.1391
21.9916 | -103.1453
21.9954 | -103.151
21.9991 | -103.1517
22.0028 | -103.1553
22.0065 | -103.1592
22.0102 | -103.1622
22.0139 | -103.1664
22.0177 | -103.1697
22.0214 | -103.1701
22.0251 | -103.1725
22.0288 | -103.1737
22.0325 | -103.175
22.0363 | -103.1759
22.04 | -103.1762
22.0437 | -103.1758
22.0474 | -103.1734
22.0511 | -103.171
22.0549 | -103.1672
22.0586 | -103.1652
22.0623 | -103.1637
22.066 | -103.1612
22.0697 | -103.1587
22.0735 | -103.1561
22.0772 | -103.1497
22.0809 | -103.1437
22.0846 | -103.1394
22.0883 | -103.1366
22.0921 | -103.1342
22.0958 | -103.1294
22.0995 | -103.1259
22.1032 | -103.1234
22.1069 | -103.1201
22.1106 | -103.1171
22.1144 | -103.1122
22.1181 | -103.1075
22.1218 | -103.1013
22.1255 | -103.0945
22.1292 | -103.0903
22.133 | -103.0865
22.1367 | -103.0824
22.1404 | -103.079
22.1441 | -103.0744
22.1478 | -103.0706
22.1516 | -103.0664
22.1553 | -103.065
22.159 | -103.0645
22.1627 | -103.0639
22.1664 | -103.0646
22.1702 | -103.0647
22.1739 | -103.0651
22.1776 | -103.065
22.1813 | -103.0645
22.185 | -103.0644
22.1887 | -103.0649
22.1925 | -103.067
22.1962 | -103.0693
22.1999 | -103.072
22.2036 | -103.0752
22.2073 | -103.0773
22.2111 | -103.0805
22.2148 | -103.0827
22.2185 | -103.0863
22.2222 | -103.0893
22.2259 | -103.089
22.2297 | -103.0879
22.2334 | -103.0847
22.2371 | -103.0834
22.2408 | -103.0843
22.2445 | -103.0845
22.2483 | -103.083
22.252 | -103.0809
22.2557 | -103.0798
22.2594 | -103.0784
22.2631 | -103.0765
22.2669 | -103.0734
22.2706 | -103.0675
22.2743 | -103.063
22.278 | -103.0572
22.2817 | -103.0526
22.2854 | -103.0484
22.2892 | -103.0433
22.2929 | -103.0368
22.2966 | -103.0292
22.3003 | -103.0216
22.304 | -103.0141
22.3078 | -103.0076
22.3115 | -102.9989
22.3152 | -102.9907
22.3189 | -102.9829
22.3226 | -102.9723
22.3264 | -102.9637
22.3301 | -102.9546
22.3338 | -102.9453
22.3375 | -102.9374
22.3412 | -102.931
22.345 | -102.9236
22.3487 | -102.9137
22.3524 | -102.9061
22.3561 | -102.8996
22.3598 | -102.8926
22.3636 | -102.8873
22.3673 | -102.8817
22.371 | -102.8779
22.3747 | -102.8751
22.3784 | -102.8715
22.3821 | -102.8676
22.3859 | -102.8652
22.3896 | -102.8643
22.3933 | -102.8667
22.397 | -102.8683
22.4007 | -102.8701
22.4045 | -102.8741
22.4082 | -102.8785
22.4119 | -102.8828
22.4156 | -102.8843
22.4193 | -102.8845
22.4231 | -102.8885
22.4268 | -102.8936
22.4305 | -102.9002
22.4342 | -102.9073
22.4379 | -102.9161
22.4417 | -102.9262
22.4454 | -102.9336
22.4491 | -102.9389
22.4528 | -102.9428
22.4565 | -102.9453
22.4603 | -102.9497
22.464 | -102.9516
22.4677 | -102.9545
22.4714 | -102.9549
22.4751 | -102.9546
22.4788 | -102.9529
22.4826 | -102.9503
22.4863 | -102.9433
22.49 | -102.933
22.4937 | -102.9243
22.4974 | -102.9153
22.5012 | -102.9044
22.5049 | -102.8924
22.5086 | -102.8786
22.5123 | -102.8646
22.516 | -102.848
22.5198 | -102.8326
22.5235 | -102.8173
22.5272 | -102.7974
22.5309 | -102.775
22.5346 | -102.7433
22.5384 | -102.7219
22.5421 | -102.6989
22.5458 | -102.6755
22.5495 | -102.651
22.5532 | -102.6272
22.557 | -102.5996
22.5607 | -102.5706
22.5644 | -102.5419
22.5681 | -102.5143
22.5718 | -102.4852
22.5755 | -102.4545
22.5793 | -102.426
22.583 | -102.3988
22.5867 | -102.3746
22.5904 | -102.3507
22.5941 | -102.324
22.5979 | -102.2989
22.6016 | -102.274
22.6053 | -102.2509
22.609 | -102.2292
22.6127 | -102.2111
22.6165 | -102.1941
22.6202 | -102.1776
22.6239 | -102.162
22.6276 | -102.1488
22.6313 | -102.1379
22.6351 | -102.1296
22.6388 | -102.1231
22.6425 | -102.1194
22.6462 | -102.1158
22.6499 | -102.113
22.6536 | -102.112
22.6574 | -102.1115
22.6611 | -102.1135
22.6648 | -102.1181
22.6685 | -102.1254
22.6722 | -102.1318
22.676 | -102.1412
22.6797 | -102.1527
22.6834 | -102.1647
22.6871 | -102.1781
22.6908 | -102.1922
22.6946 | -102.2062
22.6983 | -102.2237
22.702 | -102.2421
22.7057 | -102.2601
22.7094 | -102.2791
22.7132 | -102.2989
22.7169 | -102.317
22.7206 | -102.3336
22.7243 | -102.3516
22.728 | -102.3693
22.7318 | -102.3872
22.7355 | -102.4078
22.7392 | -102.4288
22.7429 | -102.4463
22.7466 | -102.4622
22.7503 | -102.4774
22.7541 | -102.4937
22.7578 | -102.5068
22.7615 | -102.519
22.7652 | -102.5331
22.7689 | -102.5476
22.7727 | -102.5627
22.7764 | -102.5779
22.7801 | -102.5897
22.7838 | -102.6015
22.7875 | -102.6115
22.7913 | -102.621
22.795 | -102.6284
22.7987 | -102.6362
22.8024 | -102.6428
22.8061 | -102.6518
22.8099 | -102.6613
22.8136 | -102.6689
22.8173 | -102.6739
22.821 | -102.6779
22.8247 | -102.6825
22.8285 | -102.6868
22.8322 | -102.6892
22.8359 | -102.6942
22.8396 | -102.6971
22.8433 | -102.7
22.847 | -102.7056
22.8508 | -102.7118
22.8545 | -102.7148
22.8582 | -102.7161
22.8619 | -102.7198
22.8656 | -102.7253
22.8694 | -102.7316
22.8731 | -102.7385
22.8768 | -102.7444
22.8805 | -102.7521
22.8842 | -102.76
22.888 | -102.7689
22.8917 | -102.7756
22.8954 | -102.7835
22.8991 | -102.793
22.9028 | -102.8006
22.9066 | -102.8085
22.9103 | -102.816
22.914 | -102.8243
22.9177 | -102.8326
22.9214 | -102.8419
22.9252 | -102.8522
22.9289 | -102.861
22.9326 | -102.8695
22.9363 | -102.8756
22.94 | -102.8818
22.9437 | -102.8871
22.9475 | -102.8944
22.9512 | -102.9002
22.9549 | -102.9054
22.9586 | -102.9093
22.9623 | -102.9123
22.9661 | -102.9146
22.9698 | -102.9151
22.9735 | -102.9168
22.9772 | -102.919
22.9809 | -102.9202
22.9847 | -102.9191
22.9884 | -102.9167
22.9921 | -102.9157
22.9958 | -102.9162
22.9995 | -102.9125
23.0033 | -102.9082
23.007 | -102.9058
23.0107 | -102.902
23.0144 | -102.896
23.0181 | -102.8898
23.0219 | -102.8809
23.0256 | -102.8732
23.0293 | -102.8657
23.033 | -102.8571
23.0367 | -102.8487
23.0404 | -102.8422
23.0442 | -102.8357
23.0479 | -102.8272
23.0516 | -102.8199
23.0553 | -102.8116
23.059 | -102.8014
23.0628 | -102.791
23.0665 | -102.7804
23.0702 | -102.7726
23.0739 | -102.765
23.0776 | -102.7583
23.0814 | -102.7508
23.0851 | -102.745
23.0888 | -102.7382
23.0925 | -102.7305
23.0962 | -102.7226
23.1 | -102.7161
23.1037 | -102.7075
23.1074 | -102.701
23.1111 | -102.6964
23.1148 | -102.6916
23.1185 | -102.688
23.1223 | -102.6863
23.126 | -102.685
23.1297 | -102.6831
23.1334 | -102.6819
23.1371 | -102.6812
23.1409 | -102.6816
23.1446 | -102.6846
23.1483 | -102.6871
23.152 | -102.6882
23.1557 | -102.6895
23.1595 | -102.6926
23.1632 | -102.6946
23.1669 | -102.698
23.1706 | -102.7034
23.1743 | -102.7078
23.1781 | -102.711
23.1818 | -102.7169
23.1855 | -102.7233
23.1892 | -102.7276
23.1929 | -102.7315
23.1967 | -102.7349
23.2004 | -102.7402
23.2041 | -102.7475
23.2078 | -102.753
23.2115 | -102.7573
23.2152 | -102.7617
23.219 | -102.7699
23.2227 | -102.7787
23.2264 | -102.7861
23.2301 | -102.7918
23.2338 | -102.7988
23.2376 | -102.8079
23.2413 | -102.8168
23.245 | -102.8261
23.2487 | -102.8346
23.2524 | -102.8431
23.2562 | -102.8518
23.2599 | -102.8605
23.2636 | -102.8696
23.2673 | -102.8773
23.271 | -102.8872
23.2748 | -102.897
23.2785 | -102.9064
23.2822 | -102.9158
23.2859 | -102.9248
23.2896 | -102.9324
23.2934 | -102.9387
23.2971 | -102.9498
23.3008 | -102.9602
23.3045 | -102.9689
23.3082 | -102.9785
23.3119 | -102.9864
23.3157 | -102.9955
23.3194 | -103.0039
23.3231 | -103.0121
23.3268 | -103.0223
23.3305 | -103.0306
23.3343 | -103.0387
23.338 | -103.0488
23.3417 | -103.059
23.3454 | -103.0696
23.3491 | -103.0801
23.3529 | -103.0904
23.3566 | -103.1008
23.3603 | -103.1115
23.364 | -103.1225
23.3677 | -103.1336
23.3715 | -103.1439
23.3752 | -103.1532
23.3789 | -103.1627
23.3826 | -103.1732
23.3863 | -103.1841
23.3901 | -103.1943
23.3938 | -103.2045
23.3975 | -103.2142
23.4012 | -103.2238
23.4049 | -103.233
23.4086 | -103.2426
23.4124 | -103.2512
23.4161 | -103.2606
23.4198 | -103.2706
23.4235 | -103.281
23.4272 | -103.2923
23.431 | -103.3026
23.4347 | -103.314
23.4384 | -103.325
23.4421 | -103.3345
23.4458 | -103.3419
23.4496 | -103.3493
23.4533 | -103.3563
23.457 | -103.3619
23.4607 | -103.3697
23.4644 | -103.3774
23.4682 | -103.3835
23.4719 | -103.39
23.4756 | -103.3956
23.4793 | -103.404
23.483 | -103.4151
23.4868 | -103.4254
23.4905 | -103.4333
23.4942 | -103.4416
23.4979 | -103.4494
23.5016 | -103.4557
23.5053 | -103.4626
23.5091 | -103.4727
23.5128 | -103.4821
23.5165 | -103.4896
23.5202 | -103.4955
23.5239 | -103.5036
23.5277 | -103.5147
23.5314 | -103.5241
23.5351 | -103.5305
23.5388 | -103.5396
23.5425 | -103.5508
23.5463 | -103.5603
23.55 | -103.5684
23.5537 | -103.5773
23.5574 | -103.5874
23.5611 | -103.6005
23.5649 | -103.6104
23.5686 | -103.6181
23.5723 | -103.6256
23.576 | -103.6339
23.5797 | -103.6435
23.5834 | -103.6522
23.5872 | -103.6608
23.5909 | -103.6704
23.5946 | -103.679
23.5983 | -103.689
23.602 | -103.6968
23.6058 | -103.7025
23.6095 | -103.7086
23.6132 | -103.7144
23.6169 | -103.7197
23.6206 | -103.7245
23.6244 | -103.7296
23.6281 | -103.7351
23.6318 | -103.7364
23.6355 | -103.7378
23.6392 | -103.7406
23.643 | -103.7419
23.6467 | -103.7429
23.6504 | -103.7423
23.6541 | -103.7417
23.6578 | -103.7399
23.6616 | -103.7371
23.6653 | -103.7359
23.669 | -103.7345
23.6727 | -103.7318
23.6764 | -103.7297
23.6801 | -103.728
23.6839 | -103.7251
23.6876 | -103.7217
23.6913 | -103.7178
23.695 | -103.7139
23.6987 | -103.7098
23.7025 | -103.7053
23.7062 | -103.7013
23.7099 | -103.6983
23.7136 | -103.694
23.7173 | -103.6895
23.7211 | -103.6852
23.7248 | -103.6824
23.7285 | -103.6807
23.7322 | -103.6792
23.7359 | -103.6774
23.7397 | -103.6751
23.7434 | -103.6707
23.7471 | -103.669
23.7508 | -103.6666
23.7545 | -103.6649
23.7583 | -103.6628
23.762 | -103.6604
23.7657 | -103.6596
23.7694 | -103.6614
23.7731 | -103.6629
23.7768 | -103.6655
23.7806 | -103.6692
23.7843 | -103.6722
23.788 | -103.6765
23.7917 | -103.6826
23.7954 | -103.6868
23.7992 | -103.6896
23.8029 | -103.6938
23.8066 | -103.6989
23.8103 | -103.7049
23.814 | -103.709
23.8178 | -103.713
23.8215 | -103.7209
23.8252 | -103.7295
23.8289 | -103.7368
23.8326 | -103.7449
23.8364 | -103.7536
23.8401 | -103.7621
23.8438 | -103.7689
23.8475 | -103.7751
23.8512 | -103.7822
23.855 | -103.7915
23.8587 | -103.8004
23.8624 | -103.8056
23.8661 | -103.8099
23.8698 | -103.8164
23.8735 | -103.8226
23.8773 | -103.8286
23.881 | -103.8354
23.8847 | -103.8421
23.8884 | -103.8486
23.8921 | -103.8561
23.8959 | -103.8632
23.8996 | -103.8704
23.9033 | -103.8774
23.907 | -103.8858
23.9107 | -103.894
23.9145 | -103.9
23.9182 | -103.9059
23.9219 | -103.9124
23.9256 | -103.9177
23.9293 | -103.9249
23.9331 | -103.931
23.9368 | -103.9346
23.9405 | -103.9382
23.9442 | -103.9414
23.9479 | -103.9462
23.9517 | -103.9499
23.9554 | -103.9539
23.9591 | -103.9599
23.9628 | -103.9656
23.9665 | -103.9704
23.9702 | -103.9742
23.974 | -103.9777
23.9777 | -103.9801
23.9814 | -103.9824
23.9851 | -103.9832
23.9888 | -103.9843
23.9926 | -103.9853
23.9963 | -103.9851
24 | -103.9858
24.0037 | -103.9853
24.0074 | -103.9843
24.0112 | -103.9845
24.0149 | -103.9859
24.0186 | -103.9867
24.0223 | -103.9875
24.026 | -103.9887
24.0298 | -103.9898
24.0335 | -103.9911
24.0372 | -103.9938
24.0409 | -103.9959
24.0446 | -103.9983
24.0483 | -104.0011
24.0521 | -104.0035
24.0558 | -104.0061
24.0595 | -104.0099
24.0632 | -104.0125
24.0669 | -104.0138
24.0707 | -104.0148
24.0744 | -104.0165
24.0781 | -104.0207
24.0818 | -104.0274
24.0855 | -104.0326
24.0893 | -104.0385
24.093 | -104.0447
24.0967 | -104.0512
24.1004 | -104.0567
24.1041 | -104.0618
24.1079 | -104.0661
24.1116 | -104.0704
24.1153 | -104.0748
24.119 | -104.0798
24.1227 | -104.0869
24.1265 | -104.0935
24.1302 | -104.0979
24.1339 | -104.1027
24.1376 | -104.1092
24.1413 | -104.1156
24.145 | -104.1219
24.1488 | -104.1277
24.1525 | -104.1326
24.1562 | -104.1376
24.1599 | -104.1414
24.1636 | -104.1449
24.1674 | -104.1499
24.1711 | -104.1558
24.1748 | -104.161
24.1785 | -104.1667
24.1822 | -104.1709
24.186 | -104.1747
24.1897 | -104.18
24.1934 | -104.1857
24.1971 | -104.1889
24.2008 | -104.1899
24.2046 | -104.1934
24.2083 | -104.1999
24.212 | -104.2028
24.2157 | -104.2052
24.2194 | -104.2075
24.2232 | -104.2092
24.2269 | -104.2114
24.2306 | -104.2126
24.2343 | -104.2146
24.238 | -104.2158
24.2417 | -104.2171
24.2455 | -104.2187
24.2492 | -104.2189
24.2529 | -104.2191
24.2566 | -104.2204
24.2603 | -104.2217
24.2641 | -104.223
24.2678 | -104.2246
24.2715 | -104.2241
24.2752 | -104.2207
24.2789 | -104.2192
24.2827 | -104.2182
24.2864 | -104.2155
24.2901 | -104.2115
24.2938 | -104.2102
24.2975 | -104.2097
24.3013 | -104.2064
24.305 | -104.202
24.3087 | -104.1968
24.3124 | -104.1934
24.3161 | -104.1882
24.3199 | -104.1819
24.3236 | -104.177
24.3273 | -104.1723
24.331 | -104.1659
24.3347 | -104.159
24.3384 | -104.153
24.3422 | -104.1453
24.3459 | -104.1393
24.3496 | -104.1317
24.3533 | -104.1228
24.357 | -104.1132
24.3608 | -104.1039
24.3645 | -104.0964
24.3682 | -104.0876
24.3719 | -104.0789
24.3756 | -104.0696
24.3794 | -104.0584
24.3831 | -104.0492
24.3868 | -104.0416
24.3905 | -104.0331
24.3942 | -104.0237
24.398 | -104.0137
24.4017 | -104.0027
24.4054 | -103.9923
24.4091 | -103.9828
24.4128 | -103.9744
24.4166 | -103.9676
24.4203 | -103.9604
24.424 | -103.952
24.4277 | -103.9443
24.4314 | -103.9348
24.4351 | -103.926
24.4389 | -103.9196
24.4426 | -103.9149
24.4463 | -103.9102
24.45 | -103.9055
24.4537 | -103.9006
24.4575 | -103.895
24.4612 | -103.8933
24.4649 | -103.8907
24.4686 | -103.8876
24.4723 | -103.8868
24.4761 | -103.887
24.4798 | -103.8861
24.4835 | -103.8866
24.4872 | -103.8867
24.4909 | -103.8854
24.4947 | -103.8864
24.4984 | -103.8888
24.5021 | -103.8917
24.5058 | -103.8961
24.5095 | -103.8995
24.5132 | -103.9026
24.517 | -103.907
24.5207 | -103.9096
24.5244 | -103.9133
24.5281 | -103.9193
24.5318 | -103.9256
24.5356 | -103.9304
24.5393 | -103.9347
24.543 | -103.9414
24.5467 | -103.9503
24.5504 | -103.9593
24.5542 | -103.9664
24.5579 | -103.9747
24.5616 | -103.9829
24.5653 | -103.989
24.569 | -103.992
24.5728 | -103.9941
24.5765 | -103.9987
24.5802 | -104.0009
24.5839 | -103.9984
24.5876 | -103.9971
24.5914 | -103.9904
24.5951 | -103.9704
24.5988 | -103.9441
24.6025 | -103.9249
24.6062 | -103.9033
24.6099 | -103.8737
24.6137 | -103.8426
24.6174 | -103.8165
24.6211 | -103.792
24.6248 | -103.7642
24.6285 | -103.7329
24.6323 | -103.6997
24.636 | -103.6632
24.6397 | -103.6245
24.6434 | -103.583
24.6471 | -103.54
24.6509 | -103.494
24.6546 | -103.4453
24.6583 | -103.3929
24.662 | -103.339
24.6657 | -103.2841
24.6695 | -103.2251
24.6732 | -103.1634
24.6769 | -103.1009
24.6806 | -103.0235
24.6843 | -102.9617
24.6881 | -102.9001
24.6918 | -102.8384
24.6955 | -102.7749
24.6992 | -102.7124
24.7029 | -102.6495
24.7066 | -102.5874
24.7104 | -102.527
24.7141 | -102.4673
24.7178 | -102.4093
24.7215 | -102.3547
24.7252 | -102.3017
24.729 | -102.2517
24.7327 | -102.2042
24.7364 | -102.1593
24.7401 | -102.1186
24.7438 | -102.0815
24.7476 | -102.0468
24.7513 | -102.0164
24.755 | -101.9904
24.7587 | -101.9681
24.7624 | -101.9481
24.7662 | -101.9309
24.7699 | -101.9183
24.7736 | -101.9106
24.7773 | -101.9065
24.781 | -101.9082
24.7848 | -101.9143
24.7885 | -101.9226
24.7922 | -101.9347
24.7959 | -101.9506
24.7996 | -101.9715
24.8033 | -101.9967
24.8071 | -102.0253
24.8108 | -102.0556
24.8145 | -102.0889
24.8182 | -102.126
24.8219 | -102.1656
24.8257 | -102.2074
24.8294 | -102.2521
24.8331 | -102.2994
24.8368 | -102.349
24.8405 | -102.4017
24.8443 | -102.4554
24.848 | -102.5098
24.8517 | -102.5654
24.8554 | -102.6217
24.8591 | -102.6792
24.8629 | -102.7377
24.8666 | -102.7964
24.8703 | -102.856
24.874 | -102.9142
24.8777 | -102.972
24.8815 | -103.0298
24.8852 | -103.0867
24.8889 | -103.1434
24.8926 | -103.199
24.8963 | -103.2518
24.9 | -103.3027
24.9038 | -103.352
24.9075 | -103.4002
24.9112 | -103.4469
24.9149 | -103.4935
24.9186 | -103.5423
24.9224 | -103.5888
24.9261 | -103.6313
24.9298 | -103.6708
24.9335 | -103.7106
24.9372 | -103.7457
24.941 | -103.7821
24.9447 | -103.8273
24.9484 | -103.8729
24.9521 | -103.9056
24.9558 | -103.9354
24.9596 | -103.9668
24.9633 | -103.9929
24.967 | -104.0099
24.9707 | -104.0248
24.9744 | -104.0372
24.9781 | -104.047
24.9819 | -104.0553
24.9856 | -104.0617
24.9893 | -104.0667
24.993 | -104.0718
24.9967 | -104.0762
25.0005 | -104.0815
25.0042 | -104.0877
25.0079 | -104.0929
25.0116 | -104.0971
25.0153 | -104.0999
25.0191 | -104.1042
25.0228 | -104.1098
25.0265 | -104.1134
25.0302 | -104.1158
25.0339 | -104.1176
25.0377 | -104.1206
25.0414 | -104.1223
25.0451 | -104.1231
25.0488 | -104.1253
25.0525 | -104.127
25.0563 | -104.129
25.06 | -104.1305
25.0637 | -104.1305
25.0674 | -104.1306
25.0711 | -104.1311
25.0748 | -104.1291
25.0786 | -104.1287
25.0823 | -104.1287
25.086 | -104.1294
25.0897 | -104.1298
25.0934 | -104.1299
25.0972 | -104.1301
25.1009 | -104.1327
25.1046 | -104.1356
25.1083 | -104.1371
25.112 | -104.139
25.1158 | -104.1426
25.1195 | -104.1469
25.1232 | -104.1508
25.1269 | -104.1544
25.1306 | -104.1587
25.1344 | -104.1622
25.1381 | -104.1661
25.1418 | -104.172
25.1455 | -104.178
25.1492 | -104.1832
25.153 | -104.1878
25.1567 | -104.1924
25.1604 | -104.198
25.1641 | -104.204
25.1678 | -104.2112
25.1715 | -104.2181
25.1753 | -104.2242
25.179 | -104.2318
25.1827 | -104.2408
25.1864 | -104.2488
25.1901 | -104.2565
25.1939 | -104.2627
25.1976 | -104.2684
25.2013 | -104.2756
25.205 | -104.2834
25.2087 | -104.2906
25.2125 | -104.2992
25.2162 | -104.3077
25.2199 | -104.3133
25.2236 | -104.3198
25.2273 | -104.3278
25.2311 | -104.337
25.2348 | -104.3459
25.2385 | -104.3549
25.2422 | -104.364
25.2459 | -104.3708
25.2497 | -104.3763
25.2534 | -104.3836
25.2571 | -104.3917
25.2608 | -104.4006
25.2645 | -104.4097
25.2682 | -104.4186
25.272 | -104.4282
25.2757 | -104.434
25.2794 | -104.4385
25.2831 | -104.445
25.2868 | -104.4541
25.2906 | -104.4645
25.2943 | -104.4745
25.298 | -104.4828
25.3017 | -104.4879
25.3054 | -104.4918
25.3092 | -104.4957
25.3129 | -104.4997
25.3166 | -104.5027
25.3203 | -104.5067
25.324 | -104.5098
25.3278 | -104.5086
25.3315 | -104.5089
25.3352 | -104.5107
25.3389 | -104.5126
25.3426 | -104.5156
25.3464 | -104.5162
25.3501 | -104.5169
25.3538 | -104.517
25.3575 | -104.516
25.3612 | -104.5127
25.3649 | -104.5108
25.3687 | -104.5095
25.3724 | -104.5079
25.3761 | -104.5044
25.3798 | -104.5009
25.3835 | -104.4976
25.3873 | -104.4918
25.391 | -104.4854
25.3947 | -104.4797
25.3984 | -104.4747
25.4021 | -104.4703
25.4059 | -104.4642
25.4096 | -104.458
25.4133 | -104.453
25.417 | -104.4467
25.4207 | -104.4413
25.4245 | -104.4339
25.4282 | -104.4253
25.4319 | -104.418
25.4356 | -104.4099
25.4393 | -104.4
25.443 | -104.3905
25.4468 | -104.3805
25.4505 | -104.371
25.4542 | -104.363
25.4579 | -104.3547
25.4616 | -104.3446
25.4654 | -104.3341
25.4691 | -104.3236
25.4728 | -104.3127
25.4765 | -104.3034
25.4802 | -104.2941
25.484 | -104.2858
25.4877 | -104.2785
25.4914 | -104.2713
25.4951 | -104.262
25.4988 | -104.2539
25.5026 | -104.2455
25.5063 | -104.2376
25.51 | -104.2294
25.5137 | -104.2209
25.5174 | -104.2131
25.5212 | -104.2075
25.5249 | -104.2017
25.5286 | -104.195
25.5323 | -104.1881
25.536 | -104.1822
25.5397 | -104.1749
25.5435 | -104.1672
25.5472 | -104.1608
25.5509 | -104.1554
25.5546 | -104.1503
25.5583 | -104.1444
25.5621 | -104.1369
25.5658 | -104.1311
25.5695 | -104.1248
25.5732 | -104.1183
25.5769 | -104.1124
25.5807 | -104.1075
25.5844 | -104.1042
25.5881 | -104.1009
25.5918 | -104.0961
25.5955 | -104.0906
25.5993 | -104.086
25.603 | -104.0829
25.6067 | -104.0821
25.6104 | -104.0816
25.6141 | -104.0812
25.6179 | -104.0801
25.6216 | -104.079
25.6253 | -104.0792
25.629 | -104.0806
25.6327 | -104.0813
25.6364 | -104.0839
25.6402 | -104.0869
25.6439 | -104.0898
25.6476 | -104.0935
25.6513 | -104.0982
25.655 | -104.1022
25.6588 | -104.107
25.6625 | -104.1147
25.6662 | -104.1216
25.6699 | -104.1267
25.6736 | -104.1336
25.6774 | -104.1406
25.6811 | -104.1459
25.6848 | -104.1522
25.6885 | -104.1607
25.6922 | -104.1679
25.696 | -104.175
25.6997 | -104.1818
25.7034 | -104.1872
25.7071 | -104.1953
25.7108 | -104.2017
25.7146 | -104.2067
25.7183 | -104.2118
25.722 | -104.2191
25.7257 | -104.2252
25.7294 | -104.2301
25.7331 | -104.2344
25.7369 | -104.2386
25.7406 | -104.2422
25.7443 | -104.2469
25.748 | -104.253
25.7517 | -104.2514
25.7555 | -104.2545
25.7592 | -104.2584
25.7629 | -104.2608
25.7666 | -104.2636
25.7703 | -104.2663
25.7741 | -104.2672
25.7778 | -104.2666
25.7815 | -104.267
25.7852 | -104.2656
25.7889 | -104.267
25.7927 | -104.2705
25.7964 | -104.2723
25.8001 | -104.273
25.8038 | -104.274
25.8075 | -104.2734
25.8113 | -104.2719
25.815 | -104.2689
25.8187 | -104.2662
25.8224 | -104.2616
25.8261 | -104.2579
25.8298 | -104.2559
25.8336 | -104.2528
25.8373 | -104.2502
25.841 | -104.2465
25.8447 | -104.2425
25.8484 | -104.2381
25.8522 | -104.2336
25.8559 | -104.2298
25.8596 | -104.2261
25.8633 | -104.2195
25.867 | -104.2133
25.8708 | -104.2081
25.8745 | -104.2039
25.8782 | -104.2002
25.8819 | -104.1965
25.8856 | -104.1909
25.8894 | -104.1866
25.8931 | -104.184
25.8968 | -104.1823
25.9005 | -104.1777
25.9042 | -104.1748
25.9079 | -104.1722
25.9117 | -104.1694
25.9154 | -104.1684
25.9191 | -104.1666
25.9228 | -104.1655
25.9265 | -104.1663
25.9303 | -104.1655
25.934 | -104.1638
25.9377 | -104.1639
25.9414 | -104.1646
25.9451 | -104.166
25.9489 | -104.1661
25.9526 | -104.1647
25.9563 | -104.1648
25.96 | -104.1674
25.9637 | -104.1701
25.9675 | -104.1732
25.9712 | -104.1744
25.9749 | -104.1745
25.9786 | -104.1735
25.9823 | -104.1753
25.9861 | -104.1771
25.9898 | -104.1785
25.9935 | -104.1798
25.9972 | -104.1819
26.0009 | -104.1837
26.0046 | -104.1871
26.0084 | -104.1892
26.0121 | -104.1905
26.0158 | -104.1918
26.0195 | -104.1948
26.0232 | -104.1977
26.027 | -104.1988
26.0307 | -104.1973
26.0344 | -104.198
26.0381 | -104.197
26.0418 | -104.1965
26.0456 | -104.1972
26.0493 | -104.1978
26.053 | -104.1978
26.0567 | -104.1969
26.0604 | -104.1934
26.0642 | -104.1903
26.0679 | -104.1873
26.0716 | -104.1833
26.0753 | -104.1797
26.079 | -104.1755
26.0828 | -104.1718
26.0865 | -104.1677
26.0902 | -104.1631
26.0939 | -104.1578
26.0976 | -104.1539
26.1013 | -104.1502
26.1051 | -104.1463
26.1088 | -104.1422
26.1125 | -104.1357
26.1162 | -104.1306
26.1199 | -104.1275
26.1237 | -104.1217
26.1274 | -104.1146
26.1311 | -104.1088
26.1348 | -104.1026
26.1385 | -104.0965
26.1423 | -104.092
26.146 | -104.087
26.1497 | -104.0823
26.1534 | -104.0781
26.1571 | -104.0717
26.1609 | -104.0652
26.1646 | -104.061
26.1683 | -104.0574
26.172 | -104.052
26.1757 | -104.0493
26.1795 | -104.0483
26.1832 | -104.0474
26.1869 | -104.0463
26.1906 | -104.0463
26.1943 | -104.0459
26.198 | -104.045
26.2018 | -104.0458
26.2055 | -104.0465
26.2092 | -104.0453
26.2129 | -104.0445
26.2166 | -104.0429
26.2204 | -104.041
26.2241 | -104.0409
26.2278 | -104.0411
26.2315 | -104.0427
26.2352 | -104.0441
26.239 | -104.0463
26.2427 | -104.0481
26.2464 | -104.0499
26.2501 | -104.0507
26.2538 | -104.0521
26.2576 | -104.0529
26.2613 | -104.0534
26.265 | -104.0546
26.2687 | -104.0575
26.2724 | -104.0601
26.2762 | -104.062
26.2799 | -104.0629
26.2836 | -104.0621
26.2873 | -104.063
26.291 | -104.064
26.2947 | -104.0667
26.2985 | -104.0694
26.3022 | -104.0717
26.3059 | -104.0733
26.3096 | -104.0737
26.3133 | -104.076
26.3171 | -104.0787
26.3208 | -104.0811
26.3245 | -104.0846
26.3282 | -104.0865
26.3319 | -104.0865
26.3357 | -104.0874
26.3394 | -104.0883
26.3431 | -104.0905
26.3468 | -104.094
26.3505 | -104.0976
26.3543 | -104.098
26.358 | -104.098
26.3617 | -104.0966
26.3654 | -104.0947
26.3691 | -104.0927
26.3728 | -104.0921
26.3766 | -104.0921
26.3803 | -104.0923
26.384 | -104.0927
26.3877 | -104.0924
26.3914 | -104.0913
26.3952 | -104.0908
26.3989 | -104.0907
26.4026 | -104.0947
26.4063 | -104.0979
26.41 | -104.1002
26.4138 | -104.101
26.4175 | -104.1008
26.4212 | -104.1012
26.4249 | -104.101
26.4286 | -104.1015
26.4324 | -104.1024
26.4361 | -104.1019
26.4398 | -104.1024
26.4435 | -104.1006
26.4472 | -104.0989
26.451 | -104.0995
26.4547 | -104.0986
26.4584 | -104.0968
26.4621 | -104.0949
26.4658 | -104.0935
26.4695 | -104.0886
26.4733 | -104.0841
26.477 | -104.0805
26.4807 | -104.077
26.4844 | -104.0751
26.4881 | -104.0721
26.4919 | -104.0676
26.4956 | -104.0626
26.4993 | -104.0559
26.503 | -104.0489
26.5067 | -104.0414
26.5105 | -104.0352
26.5142 | -104.0295
26.5179 | -104.0251
26.5216 | -104.0199
26.5253 | -104.0142
26.5291 | -104.0086
26.5328 | -104
26.5365 | -103.9914
26.5402 | -103.9838
26.5439 | -103.9768
26.5477 | -103.9685
26.5514 | -103.9603
26.5551 | -103.953
26.5588 | -103.9449
26.5625 | -103.9366
26.5662 | -103.9291
26.57 | -103.9212
26.5737 | -103.9126
26.5774 | -103.9026
26.5811 | -103.8953
26.5848 | -103.8867
26.5886 | -103.8781
26.5923 | -103.8708
26.596 | -103.8617
26.5997 | -103.8539
26.6034 | -103.8469
26.6072 | -103.8393
26.6109 | -103.8311
26.6146 | -103.8228
26.6183 | -103.8127
26.622 | -103.8041
26.6258 | -103.7961
26.6295 | -103.7865
26.6332 | -103.7778
26.6369 | -103.7697
26.6406 | -103.7619
26.6444 | -103.7545
26.6481 | -103.749
26.6518 | -103.7441
26.6555 | -103.7397
26.6592 | -103.7337
26.6629 | -103.7271
26.6667 | -103.719
26.6704 | -103.7107
26.6741 | -103.7025
26.6778 | -103.6954
26.6815 | -103.6892
26.6853 | -103.6838
26.689 | -103.6753
26.6927 | -103.6671
26.6964 | -103.6607
26.7001 | -103.6539
26.7039 | -103.6484
26.7076 | -103.6451
26.7113 | -103.6411
26.715 | -103.6372
26.7187 | -103.6335
26.7225 | -103.6296
26.7262 | -103.627
26.7299 | -103.6256
26.7336 | -103.6253
26.7373 | -103.6249
26.7411 | -103.6257
26.7448 | -103.6269
26.7485 | -103.6258
26.7522 | -103.6245
26.7559 | -103.6236
26.7596 | -103.6233
26.7634 | -103.6247
26.7671 | -103.6256
26.7708 | -103.6256
26.7745 | -103.626
26.7782 | -103.6256
26.782 | -103.6248
26.7857 | -103.6252
26.7894 | -103.6271
26.7931 | -103.63
26.7968 | -103.6322
26.8006 | -103.6361
26.8043 | -103.6405
26.808 | -103.6463
26.8117 | -103.6517
26.8154 | -103.6567
26.8192 | -103.6625
26.8229 | -103.6668
26.8266 | -103.6714
26.8303 | -103.6743
26.834 | -103.678
26.8377 | -103.6815
26.8415 | -103.6861
26.8452 | -103.6914
26.8489 | -103.6992
26.8526 | -103.7052
26.8563 | -103.7086
26.8601 | -103.7113
26.8638 | -103.7143
26.8675 | -103.7175
26.8712 | -103.7218
26.8749 | -103.7273
26.8787 | -103.7326
26.8824 | -103.7366
26.8861 | -103.7411
26.8898 | -103.745
26.8935 | -103.747
26.8973 | -103.7479
26.901 | -103.7497
26.9047 | -103.7523
26.9084 | -103.7541
26.9121 | -103.7559
26.9159 | -103.7564
26.9196 | -103.758
26.9233 | -103.7609
26.927 | -103.7634
26.9307 | -103.7655
26.9344 | -103.7668
26.9382 | -103.7682
26.9419 | -103.7714
26.9456 | -103.7739
26.9493 | -103.7755
26.953 | -103.7771
26.9568 | -103.7815
26.9605 | -103.7852
26.9642 | -103.7899
26.9679 | -103.7932
26.9716 | -103.7956
26.9754 | -103.7982
26.9791 | -103.8006
26.9828 | -103.8031
26.9865 | -103.8064
26.9902 | -103.8091
26.994 | -103.8118
26.9977 | -103.8159
27.0014 | -103.8189
27.0051 | -103.8217
27.0088 | -103.8259
27.0126 | -103.8305
27.0163 | -103.8339
27.02 | -103.8362
27.0237 | -103.838
27.0274 | -103.8386
27.0311 | -103.8382
27.0349 | -103.836
27.0386 | -103.8351
27.0423 | -103.8357
27.046 | -103.8348
27.0497 | -103.8347
27.0535 | -103.8358
27.0572 | -103.8351
27.0609 | -103.8344
27.0646 | -103.8327
27.0683 | -103.8304
27.0721 | -103.828
27.0758 | -103.8257
27.0795 | -103.824
27.0832 | -103.8212
27.0869 | -103.816
27.0907 | -103.8091
27.0944 | -103.8015
27.0981 | -103.7941
27.1018 | -103.7881
27.1055 | -103.7809
27.1093 | -103.7732
27.113 | -103.7651
27.1167 | -103.757
27.1204 | -103.7484
27.1241 | -103.7395
27.1278 | -103.7301
27.1316 | -103.7191
27.1353 | -103.7069
27.139 | -103.6948
27.1427 | -103.6833
27.1464 | -103.6735
27.1502 | -103.6646
27.1539 | -103.6547
27.1576 | -103.6425
27.1613 | -103.6314
27.165 | -103.6193
27.1688 | -103.6064
27.1725 | -103.5941
27.1762 | -103.5808
27.1799 | -103.5668
27.1836 | -103.5528
27.1874 | -103.5379
27.1911 | -103.5233
27.1948 | -103.5091
27.1985 | -103.4964
27.2022 | -103.4854
27.206 | -103.4743
27.2097 | -103.4633
27.2134 | -103.4529
27.2171 | -103.4444
27.2208 | -103.4357
27.2245 | -103.4252
27.2283 | -103.415
27.232 | -103.4045
27.2357 | -103.3946
27.2394 | -103.3853
27.2431 | -103.3775
27.2469 | -103.3716
27.2506 | -103.3664
27.2543 | -103.3606
27.258 | -103.3541
27.2617 | -103.3481
27.2655 | -103.3439
27.2692 | -103.3398
27.2729 | -103.337
27.2766 | -103.3346
27.2803 | -103.332
27.2841 | -103.3306
27.2878 | -103.33
27.2915 | -103.3313
27.2952 | -103.3332
27.2989 | -103.3347
27.3026 | -103.3341
27.3064 | -103.3322
27.3101 | -103.3321
27.3138 | -103.3323
27.3175 | -103.3336
27.3212 | -103.335
27.325 | -103.3356
27.3287 | -103.3376
27.3324 | -103.3417
27.3361 | -103.3471
27.3398 | -103.3498
27.3436 | -103.3523
27.3473 | -103.355
27.351 | -103.359
27.3547 | -103.3624
27.3584 | -103.3653
27.3622 | -103.3689
27.3659 | -103.3713
27.3696 | -103.3737
27.3733 | -103.3761
27.377 | -103.3788
27.3808 | -103.3815
27.3845 | -103.3853
27.3882 | -103.3892
27.3919 | -103.3919
27.3956 | -103.3937
27.3993 | -103.3958
27.4031 | -103.3993
27.4068 | -103.4026
27.4105 | -103.4066
27.4142 | -103.4103
27.4179 | -103.4156
27.4217 | -103.4223
27.4254 | -103.4276
27.4291 | -103.4315
27.4328 | -103.4369
27.4365 | -103.441
27.4403 | -103.4439
27.444 | -103.4464
27.4477 | -103.4487
27.4514 | -103.454
27.4551 | -103.4598
27.4589 | -103.4629
27.4626 | -103.465
27.4663 | -103.4681
27.47 | -103.4718
27.4737 | -103.4761
27.4775 | -103.4831
27.4812 | -103.4883
27.4849 | -103.4907
27.4886 | -103.4907
27.4923 | -103.4926
27.496 | -103.4965
27.4998 | -103.5004
27.5035 | -103.5048
27.5072 | -103.5081
27.5109 | -103.511
27.5146 | -103.515
27.5184 | -103.5212
27.5221 | -103.5282
27.5258 | -103.5329
27.5295 | -103.5357
27.5332 | -103.5396
27.537 | -103.5419
27.5407 | -103.5455
27.5444 | -103.5525
27.5481 | -103.5571
27.5518 | -103.5622
27.5556 | -103.5692
27.5593 | -103.5736
27.563 | -103.5788
27.5667 | -103.5827
27.5704 | -103.5861
27.5742 | -103.5895
27.5779 | -103.5938
27.5816 | -103.5979
27.5853 | -103.6009
27.589 | -103.6037
27.5927 | -103.6071
27.5965 | -103.61
27.6002 | -103.6137
27.6039 | -103.6161
27.6076 | -103.6194
27.6113 | -103.6232
27.6151 | -103.625
27.6188 | -103.6279
27.6225 | -103.6313
27.6262 | -103.6356
27.6299 | -103.639
27.6337 | -103.6389
27.6374 | -103.639
27.6411 | -103.6411
27.6448 | -103.6424
27.6485 | -103.6439
27.6523 | -103.6451
27.656 | -103.6447
27.6597 | -103.6439
27.6634 | -103.6439
27.6671 | -103.6445
27.6709 | -103.6444
27.6746 | -103.6429
27.6783 | -103.6408
27.682 | -103.6403
27.6857 | -103.6394
27.6894 | -103.6418
27.6932 | -103.6446
27.6969 | -103.6436
27.7006 | -103.6436
27.7043 | -103.6435
27.708 | -103.6435
27.7118 | -103.6445
27.7155 | -103.645
27.7192 | -103.644
27.7229 | -103.6426
27.7266 | -103.6426
27.7304 | -103.6433
27.7341 | -103.6442
27.7378 | -103.6463
27.7415 | -103.6484
27.7452 | -103.6505
27.749 | -103.6517
27.7527 | -103.6552
27.7564 | -103.658
27.7601 | -103.661
27.7638 | -103.6639
27.7675 | -103.6684
27.7713 | -103.6729
27.775 | -103.6741
27.7787 | -103.6759
27.7824 | -103.6789
27.7861 | -103.6835
27.7899 | -103.6875
27.7936 | -103.6896
27.7973 | -103.6922
27.801 | -103.6929
27.8047 | -103.6946
27.8085 | -103.6966
27.8122 | -103.6992
27.8159 | -103.7019
27.8196 | -103.7045
27.8233 | -103.7077
27.8271 | -103.7105
27.8308 | -103.713
27.8345 | -103.7164
27.8382 | -103.7199
27.8419 | -103.7213
27.8457 | -103.7233
27.8494 | -103.7253
27.8531 | -103.7268
27.8568 | -103.727
27.8605 | -103.7275
27.8642 | -103.7283
27.868 | -103.7277
27.8717 | -103.7278
27.8754 | -103.7289
27.8791 | -103.7303
27.8828 | -103.7288
27.8866 | -103.7258
27.8903 | -103.7223
27.894 | -103.7194
27.8977 | -103.7181
27.9014 | -103.7171
27.9052 | -103.7164
27.9089 | -103.7152
27.9126 | -103.7127
27.9163 | -103.7103
27.92 | -103.71
27.9238 | -103.7096
27.9275 | -103.7068
27.9312 | -103.7043
27.9349 | -103.7014
27.9386 | -103.698
27.9424 | -103.696
27.9461 | -103.6922
27.9498 | -103.6877
27.9535 | -103.6853
27.9572 | -103.6841
27.9609 | -103.6843
27.9647 | -103.685
27.9684 | -103.6865
27.9721 | -103.687
27.9758 | -103.6872
27.9795 | -103.6878
27.9833 | -103.6886
27.987 | -103.6882
27.9907 | -103.6886
27.9944 | -103.6893
27.9981 | -103.6901
28.0019 | -103.6901