## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 230 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 232
Hood | 02-92B/P | 171
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1015.1 | hPa |  | PRC PRO 01
Humidity | 48.1 | %/Hr
Internal temperature of calibration bench | 23.75 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1028 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.1 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.9 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.03 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3120 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 17/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 24/05/2018
Engineer | Anthony HUE | Approval |  | 29/05/2018
CEO | Robert GLEMOT | Approval |  | 29/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3284 | 57.7713
0.012209 | -36.0176 | 38.8371
0.018313 | -35.0378 | 28.4261
0.024417 | -34.5832 | 21.9791
0.030521 | -34.3987 | 17.9451
0.036626 | -34.2489 | 14.9169
0.04273 | -34.1547 | 12.7607
0.048834 | -34.1033 | 11.1103
0.054938 | -34.0734 | 9.8546
0.061043 | -34.0475 | 8.8641
0.067147 | -34.0321 | 8.0572
0.073251 | -34.0177 | 7.311
0.079355 | -34.0066 | 6.6671
0.08546 | -33.9978 | 6.1198
0.091564 | -33.9922 | 5.6592
0.097668 | -33.9903 | 5.2651
0.10377 | -33.9894 | 4.9432
0.10988 | -33.9876 | 4.6314
0.11598 | -33.9854 | 4.3484
0.12209 | -33.9832 | 4.0865
0.12819 | -33.9811 | 3.8434
0.13429 | -33.9793 | 3.6208
0.1404 | -33.9777 | 3.4354
0.1465 | -33.9761 | 3.2531
0.15261 | -33.9751 | 3.0868
0.15871 | -33.9749 | 2.9334
0.16482 | -33.9756 | 2.7892
0.17092 | -33.9772 | 2.658
0.17702 | -33.9782 | 2.5282
0.18313 | -33.9792 | 2.4064
0.18923 | -33.9798 | 2.2926
0.19534 | -33.9801 | 2.1864
0.20144 | -33.9801 | 2.0867
0.20754 | -33.9804 | 1.9981
0.21365 | -33.9806 | 1.9082
0.21975 | -33.981 | 1.8227
0.22586 | -33.9814 | 1.741
0.23196 | -33.9819 | 1.6626
0.23807 | -33.9825 | 1.5867
0.24417 | -33.983 | 1.5164
0.25027 | -33.9836 | 1.4455
0.25638 | -33.9842 | 1.3773
0.26248 | -33.9848 | 1.3117
0.26859 | -33.9855 | 1.2488
0.27469 | -33.9861 | 1.1885
0.2808 | -33.9867 | 1.1345
0.2869 | -33.9873 | 1.0793
0.293 | -33.9879 | 1.0262
0.29911 | -33.9885 | 0.97519
0.30521 | -33.9891 | 0.92583
0.31132 | -33.9898 | 0.88032
0.31742 | -33.9904 | 0.83384
0.32353 | -33.9911 | 0.78861
0.32963 | -33.9918 | 0.74456
0.33573 | -33.9925 | 0.70147
0.34184 | -33.9932 | 0.65938
0.34794 | -33.9938 | 0.61993
0.35405 | -33.9945 | 0.57973
0.36015 | -33.9951 | 0.54052
0.36626 | -33.9957 | 0.50218
0.37236 | -33.9963 | 0.46496
0.37846 | -33.9969 | 0.42863
0.38457 | -33.9973 | 0.39513
0.39067 | -33.9978 | 0.36062
0.39678 | -33.9983 | 0.32714
0.40288 | -33.9988 | 0.29449
0.40899 | -33.9993 | 0.26241
0.41509 | -33.9997 | 0.23108
0.42119 | -34.0002 | 0.20131
0.4273 | -34.0006 | 0.17097
0.4334 | -34.001 | 0.14117
0.43951 | -34.0015 | 0.11178
0.44561 | -34.002 | 0.082808
0.45172 | -34.0025 | 0.055007
0.45782 | -34.0029 | 0.026899
0.46392 | -34.0034 | -0.0008743
0.47003 | -34.0039 | -0.028197
0.47613 | -34.0045 | -0.055018
0.48224 | -34.005 | -0.081392
0.48834 | -34.0056 | -0.10649
0.49445 | -34.0062 | -0.13208
0.50055 | -34.0068 | -0.15743
0.50665 | -34.0073 | -0.18236
0.51276 | -34.0079 | -0.20699
0.51886 | -34.0085 | -0.23114
0.52497 | -34.009 | -0.25444
0.53107 | -34.0095 | -0.27812
0.53717 | -34.01 | -0.30153
0.54328 | -34.0105 | -0.32477
0.54938 | -34.011 | -0.34762
0.55549 | -34.0114 | -0.37006
0.56159 | -34.0118 | -0.39169
0.5677 | -34.0122 | -0.41383
0.5738 | -34.0126 | -0.4356
0.5799 | -34.013 | -0.45715
0.58601 | -34.0133 | -0.47853
0.59211 | -34.0137 | -0.49902
0.59822 | -34.014 | -0.51991
0.60432 | -34.0144 | -0.54056
0.61043 | -34.0147 | -0.56118
0.61653 | -34.0151 | -0.58145
0.62263 | -34.0154 | -0.60168
0.62874 | -34.0158 | -0.62138
0.63484 | -34.0162 | -0.6414
0.64095 | -34.0166 | -0.6613
0.64705 | -34.017 | -0.68106
0.65316 | -34.0174 | -0.70062
0.65926 | -34.0178 | -0.72006
0.66536 | -34.0183 | -0.73903
0.67147 | -34.0187 | -0.75843
0.67757 | -34.0191 | -0.77756
0.68368 | -34.0196 | -0.79651
0.68978 | -34.02 | -0.81542
0.69589 | -34.0205 | -0.83403
0.70199 | -34.021 | -0.8519
0.70809 | -34.0214 | -0.87015
0.7142 | -34.0219 | -0.88816
0.7203 | -34.0223 | -0.9062
0.72641 | -34.0227 | -0.92384
0.73251 | -34.0232 | -0.9414
0.73862 | -34.0236 | -0.95803
0.74472 | -34.024 | -0.97518
0.75082 | -34.0244 | -0.99217
0.75693 | -34.0247 | -1.0089
0.76303 | -34.0251 | -1.0255
0.76914 | -34.0254 | -1.0412
0.77524 | -34.0257 | -1.0576
0.78135 | -34.0261 | -1.0737
0.78745 | -34.0264 | -1.0897
0.79355 | -34.0267 | -1.1057
0.79966 | -34.0271 | -1.1215
0.80576 | -34.0274 | -1.1367
0.81187 | -34.0277 | -1.1523
0.81797 | -34.028 | -1.1677
0.82408 | -34.0284 | -1.1831
0.83018 | -34.0287 | -1.1984
0.83628 | -34.0291 | -1.2138
0.84239 | -34.0294 | -1.229
0.84849 | -34.0298 | -1.2443
0.8546 | -34.0301 | -1.2596
0.8607 | -34.0305 | -1.2747
0.86681 | -34.0309 | -1.29
0.87291 | -34.0312 | -1.3052
0.87901 | -34.0316 | -1.3204
0.88512 | -34.032 | -1.3355
0.89122 | -34.0324 | -1.3506
0.89733 | -34.0327 | -1.3659
0.90343 | -34.0331 | -1.381
0.90953 | -34.0335 | -1.3961
0.91564 | -34.0338 | -1.4112
0.92174 | -34.0342 | -1.4263
0.92785 | -34.0346 | -1.4414
0.93395 | -34.0349 | -1.4565
0.94006 | -34.0353 | -1.4715
0.94616 | -34.0356 | -1.4862
0.95226 | -34.036 | -1.5011
0.95837 | -34.0363 | -1.516
0.96447 | -34.0367 | -1.5307
0.97058 | -34.037 | -1.5455
0.97668 | -34.0374 | -1.5602
0.98279 | -34.0378 | -1.5747
0.98889 | -34.0381 | -1.5894
0.99499 | -34.0384 | -1.6039
1.0011 | -34.0388 | -1.6185
1.0072 | -34.0391 | -1.633
1.0133 | -34.0395 | -1.6475
1.0194 | -34.0398 | -1.6616
1.0255 | -34.0401 | -1.676
1.0316 | -34.0405 | -1.6904
1.0377 | -34.0408 | -1.7045
1.0438 | -34.0412 | -1.7187
1.0499 | -34.0415 | -1.7326
1.056 | -34.0418 | -1.7467
1.0621 | -34.0422 | -1.7607
1.0682 | -34.0425 | -1.7747
1.0743 | -34.0428 | -1.7886
1.0805 | -34.0432 | -1.8024
1.0866 | -34.0435 | -1.816
1.0927 | -34.0438 | -1.8297
1.0988 | -34.0442 | -1.8435
1.1049 | -34.0445 | -1.8572
1.111 | -34.0448 | -1.8708
1.1171 | -34.0452 | -1.8843
1.1232 | -34.0455 | -1.8975
1.1293 | -34.0458 | -1.9109
1.1354 | -34.0461 | -1.9242
1.1415 | -34.0465 | -1.9375
1.1476 | -34.0468 | -1.9509
1.1537 | -34.0471 | -1.9642
1.1598 | -34.0474 | -1.9771
1.1659 | -34.0477 | -1.9902
1.172 | -34.048 | -2.0034
1.1781 | -34.0483 | -2.0166
1.1842 | -34.0487 | -2.0297
1.1903 | -34.049 | -2.0427
1.1964 | -34.0493 | -2.0558
1.2025 | -34.0496 | -2.0687
1.2086 | -34.0499 | -2.0816
1.2147 | -34.0502 | -2.0946
1.2209 | -34.0505 | -2.1076
1.227 | -34.0507 | -2.1204
1.2331 | -34.051 | -2.1334
1.2392 | -34.0513 | -2.1465
1.2453 | -34.0516 | -2.1594
1.2514 | -34.0519 | -2.1724
1.2575 | -34.0522 | -2.1854
1.2636 | -34.0525 | -2.1984
1.2697 | -34.0527 | -2.2114
1.2758 | -34.053 | -2.2243
1.2819 | -34.0533 | -2.2373
1.288 | -34.0535 | -2.2503
1.2941 | -34.0538 | -2.2632
1.3002 | -34.0541 | -2.2761
1.3063 | -34.0543 | -2.2891
1.3124 | -34.0546 | -2.3021
1.3185 | -34.0548 | -2.3151
1.3246 | -34.0551 | -2.3281
1.3307 | -34.0553 | -2.341
1.3368 | -34.0556 | -2.3538
1.3429 | -34.0558 | -2.3667
1.349 | -34.056 | -2.3796
1.3551 | -34.0563 | -2.3924
1.3613 | -34.0565 | -2.4054
1.3674 | -34.0568 | -2.4181
1.3735 | -34.057 | -2.4309
1.3796 | -34.0572 | -2.4437
1.3857 | -34.0575 | -2.4565
1.3918 | -34.0577 | -2.4692
1.3979 | -34.0579 | -2.4818
1.404 | -34.0582 | -2.4943
1.4101 | -34.0584 | -2.507
1.4162 | -34.0586 | -2.5196
1.4223 | -34.0588 | -2.5321
1.4284 | -34.0591 | -2.5446
1.4345 | -34.0593 | -2.5571
1.4406 | -34.0595 | -2.5692
1.4467 | -34.0597 | -2.5816
1.4528 | -34.06 | -2.5939
1.4589 | -34.0602 | -2.6062
1.465 | -34.0604 | -2.6185
1.4711 | -34.0606 | -2.6307
1.4772 | -34.0608 | -2.6426
1.4833 | -34.061 | -2.6548
1.4894 | -34.0612 | -2.6668
1.4955 | -34.0615 | -2.6789
1.5016 | -34.0617 | -2.6909
1.5078 | -34.0619 | -2.7025
1.5139 | -34.0621 | -2.7144
1.52 | -34.0623 | -2.7263
1.5261 | -34.0626 | -2.7381
1.5322 | -34.0628 | -2.75
1.5383 | -34.063 | -2.7618
1.5444 | -34.0632 | -2.7733
1.5505 | -34.0634 | -2.7851
1.5566 | -34.0637 | -2.7968
1.5627 | -34.0639 | -2.8084
1.5688 | -34.0641 | -2.82
1.5749 | -34.0643 | -2.8316
1.581 | -34.0646 | -2.8431
1.5871 | -34.0648 | -2.8547
1.5932 | -34.0651 | -2.8663
1.5993 | -34.0653 | -2.8778
1.6054 | -34.0656 | -2.8892
1.6115 | -34.0658 | -2.9007
1.6176 | -34.0661 | -2.9121
1.6237 | -34.0663 | -2.9237
1.6298 | -34.0666 | -2.9351
1.6359 | -34.0668 | -2.9466
1.642 | -34.0671 | -2.9579
1.6482 | -34.0674 | -2.9693
1.6543 | -34.0676 | -2.9807
1.6604 | -34.0679 | -2.9922
1.6665 | -34.0682 | -3.0037
1.6726 | -34.0684 | -3.0152
1.6787 | -34.0687 | -3.0265
1.6848 | -34.069 | -3.038
1.6909 | -34.0692 | -3.0493
1.697 | -34.0695 | -3.0608
1.7031 | -34.0697 | -3.0723
1.7092 | -34.07 | -3.0837
1.7153 | -34.0703 | -3.0952
1.7214 | -34.0706 | -3.1069
1.7275 | -34.0708 | -3.1185
1.7336 | -34.0711 | -3.1301
1.7397 | -34.0714 | -3.1415
1.7458 | -34.0717 | -3.153
1.7519 | -34.0719 | -3.1645
1.758 | -34.0722 | -3.176
1.7641 | -34.0724 | -3.1876
1.7702 | -34.0727 | -3.1991
1.7763 | -34.073 | -3.2106
1.7824 | -34.0732 | -3.2221
1.7885 | -34.0735 | -3.2338
1.7947 | -34.0737 | -3.2453
1.8008 | -34.0739 | -3.2569
1.8069 | -34.0742 | -3.2684
1.813 | -34.0744 | -3.28
1.8191 | -34.0747 | -3.2917
1.8252 | -34.0749 | -3.3033
1.8313 | -34.0752 | -3.3149
1.8374 | -34.0754 | -3.3265
1.8435 | -34.0756 | -3.3381
1.8496 | -34.0759 | -3.3498
1.8557 | -34.0761 | -3.3614
1.8618 | -34.0763 | -3.373
1.8679 | -34.0765 | -3.3848
1.874 | -34.0767 | -3.3965
1.8801 | -34.0769 | -3.4081
1.8862 | -34.0771 | -3.4197
1.8923 | -34.0773 | -3.4314
1.8984 | -34.0775 | -3.4432
1.9045 | -34.0777 | -3.455
1.9106 | -34.0779 | -3.4667
1.9167 | -34.0781 | -3.4784
1.9228 | -34.0783 | -3.4902
1.9289 | -34.0785 | -3.502
1.9351 | -34.0787 | -3.5138
1.9412 | -34.0789 | -3.5255
1.9473 | -34.0791 | -3.5372
1.9534 | -34.0793 | -3.5489
1.9595 | -34.0795 | -3.5606
1.9656 | -34.0797 | -3.5724
1.9717 | -34.0799 | -3.584
1.9778 | -34.08 | -3.5957
1.9839 | -34.0802 | -3.6075
1.99 | -34.0804 | -3.6192
1.9961 | -34.0806 | -3.631
2.0022 | -34.0808 | -3.6427
2.0083 | -34.081 | -3.6544
2.0144 | -34.0812 | -3.666
2.0205 | -34.0814 | -3.6776
2.0266 | -34.0815 | -3.6893
2.0327 | -34.0817 | -3.7011
2.0388 | -34.0819 | -3.7126
2.0449 | -34.0821 | -3.7243
2.051 | -34.0823 | -3.736
2.0571 | -34.0825 | -3.7475
2.0632 | -34.0827 | -3.7591
2.0693 | -34.0829 | -3.7708
2.0754 | -34.0831 | -3.7823
2.0816 | -34.0834 | -3.7938
2.0877 | -34.0836 | -3.8055
2.0938 | -34.0838 | -3.817
2.0999 | -34.084 | -3.8286
2.106 | -34.0842 | -3.8402
2.1121 | -34.0844 | -3.8516
2.1182 | -34.0847 | -3.8632
2.1243 | -34.0849 | -3.8747
2.1304 | -34.0851 | -3.8863
2.1365 | -34.0853 | -3.8978
2.1426 | -34.0855 | -3.9091
2.1487 | -34.0857 | -3.9206
2.1548 | -34.086 | -3.9319
2.1609 | -34.0862 | -3.9434
2.167 | -34.0864 | -3.9547
2.1731 | -34.0866 | -3.9661
2.1792 | -34.0869 | -3.9775
2.1853 | -34.0871 | -3.9888
2.1914 | -34.0873 | -4.0002
2.1975 | -34.0875 | -4.0116
2.2036 | -34.0878 | -4.023
2.2097 | -34.088 | -4.0344
2.2158 | -34.0883 | -4.0457
2.222 | -34.0885 | -4.0571
2.2281 | -34.0887 | -4.0685
2.2342 | -34.089 | -4.0798
2.2403 | -34.0892 | -4.0911
2.2464 | -34.0894 | -4.1023
2.2525 | -34.0897 | -4.1136
2.2586 | -34.0899 | -4.1249
2.2647 | -34.0901 | -4.1363
2.2708 | -34.0903 | -4.1475
2.2769 | -34.0906 | -4.1588
2.283 | -34.0908 | -4.1701
2.2891 | -34.0911 | -4.1814
2.2952 | -34.0913 | -4.1928
2.3013 | -34.0915 | -4.2041
2.3074 | -34.0917 | -4.2154
2.3135 | -34.092 | -4.2266
2.3196 | -34.0922 | -4.2378
2.3257 | -34.0925 | -4.2492
2.3318 | -34.0927 | -4.2605
2.3379 | -34.0929 | -4.2718
2.344 | -34.0932 | -4.2831
2.3501 | -34.0934 | -4.2943
2.3562 | -34.0936 | -4.3054
2.3623 | -34.0939 | -4.3168
2.3685 | -34.0941 | -4.328
2.3746 | -34.0943 | -4.3393
2.3807 | -34.0946 | -4.3506
2.3868 | -34.0948 | -4.3618
2.3929 | -34.095 | -4.373
2.399 | -34.0952 | -4.3844
2.4051 | -34.0954 | -4.3956
2.4112 | -34.0957 | -4.4068
2.4173 | -34.0959 | -4.418
2.4234 | -34.0961 | -4.4293
2.4295 | -34.0963 | -4.4404
2.4356 | -34.0966 | -4.4517
2.4417 | -34.0968 | -4.463
2.4478 | -34.097 | -4.4742
2.4539 | -34.0972 | -4.4854
2.46 | -34.0974 | -4.4966
2.4661 | -34.0977 | -4.5077
2.4722 | -34.0979 | -4.5189
2.4783 | -34.0981 | -4.5301
2.4844 | -34.0983 | -4.5414
2.4905 | -34.0985 | -4.5527
2.4966 | -34.0987 | -4.5639
2.5027 | -34.0989 | -4.5751
2.5089 | -34.0991 | -4.5863
2.515 | -34.0993 | -4.5976
2.5211 | -34.0995 | -4.6087
2.5272 | -34.0997 | -4.6199
2.5333 | -34.0999 | -4.6311
2.5394 | -34.1001 | -4.6422
2.5455 | -34.1003 | -4.6534
2.5516 | -34.1005 | -4.6645
2.5577 | -34.1007 | -4.6756
2.5638 | -34.1009 | -4.6866
2.5699 | -34.1012 | -4.6978
2.576 | -34.1014 | -4.7089
2.5821 | -34.1016 | -4.7201
2.5882 | -34.1018 | -4.7312
2.5943 | -34.102 | -4.7423
2.6004 | -34.1021 | -4.7534
2.6065 | -34.1023 | -4.7646
2.6126 | -34.1025 | -4.7757
2.6187 | -34.1027 | -4.7868
2.6248 | -34.1029 | -4.7978
2.6309 | -34.1031 | -4.8089
2.637 | -34.1033 | -4.8198
2.6431 | -34.1035 | -4.8308
2.6492 | -34.1037 | -4.8419
2.6554 | -34.1039 | -4.853
2.6615 | -34.1041 | -4.8641
2.6676 | -34.1043 | -4.875
2.6737 | -34.1045 | -4.8861
2.6798 | -34.1047 | -4.8969
2.6859 | -34.1049 | -4.9079
2.692 | -34.1051 | -4.919
2.6981 | -34.1052 | -4.93
2.7042 | -34.1054 | -4.941
2.7103 | -34.1056 | -4.952
2.7164 | -34.1058 | -4.963
2.7225 | -34.106 | -4.974
2.7286 | -34.1062 | -4.9849
2.7347 | -34.1064 | -4.9959
2.7408 | -34.1066 | -5.0068
2.7469 | -34.1068 | -5.0178
2.753 | -34.1069 | -5.0287
2.7591 | -34.1071 | -5.0396
2.7652 | -34.1073 | -5.0505
2.7713 | -34.1075 | -5.0614
2.7774 | -34.1077 | -5.0723
2.7835 | -34.1079 | -5.0832
2.7896 | -34.1081 | -5.0941
2.7958 | -34.1083 | -5.105
2.8019 | -34.1085 | -5.1159
2.808 | -34.1087 | -5.1268
2.8141 | -34.1089 | -5.1375
2.8202 | -34.1091 | -5.1484
2.8263 | -34.1093 | -5.1591
2.8324 | -34.1095 | -5.1699
2.8385 | -34.1097 | -5.1807
2.8446 | -34.1099 | -5.1914
2.8507 | -34.1101 | -5.2022
2.8568 | -34.1103 | -5.213
2.8629 | -34.1105 | -5.2239
2.869 | -34.1107 | -5.2346
2.8751 | -34.1109 | -5.2454
2.8812 | -34.1111 | -5.256
2.8873 | -34.1113 | -5.2668
2.8934 | -34.1115 | -5.2776
2.8995 | -34.1117 | -5.2883
2.9056 | -34.1119 | -5.2992
2.9117 | -34.1121 | -5.3099
2.9178 | -34.1124 | -5.3205
2.9239 | -34.1126 | -5.3312
2.93 | -34.1128 | -5.3419
2.9361 | -34.113 | -5.3527
2.9423 | -34.1132 | -5.3635
2.9484 | -34.1135 | -5.3741
2.9545 | -34.1137 | -5.3847
2.9606 | -34.1139 | -5.3954
2.9667 | -34.1142 | -5.4062
2.9728 | -34.1144 | -5.4168
2.9789 | -34.1146 | -5.4275
2.985 | -34.1149 | -5.4381
2.9911 | -34.1151 | -5.4488
2.9972 | -34.1153 | -5.4594
3.0033 | -34.1156 | -5.4701
3.0094 | -34.1158 | -5.4807
3.0155 | -34.116 | -5.4914
3.0216 | -34.1163 | -5.502
3.0277 | -34.1165 | -5.5126
3.0338 | -34.1167 | -5.5233
3.0399 | -34.117 | -5.5339
3.046 | -34.1172 | -5.5445
3.0521 | -34.1175 | -5.5553
3.0582 | -34.1177 | -5.566
3.0643 | -34.118 | -5.5767
3.0704 | -34.1182 | -5.5874
3.0765 | -34.1185 | -5.5981
3.0827 | -34.1187 | -5.6088
3.0888 | -34.119 | -5.6195
3.0949 | -34.1192 | -5.6302
3.101 | -34.1195 | -5.6409
3.1071 | -34.1197 | -5.6516
3.1132 | -34.1199 | -5.6623
3.1193 | -34.1202 | -5.673
3.1254 | -34.1205 | -5.6836
3.1315 | -34.1207 | -5.6943
3.1376 | -34.121 | -5.705
3.1437 | -34.1212 | -5.7157
3.1498 | -34.1215 | -5.7264
3.1559 | -34.1217 | -5.7371
3.162 | -34.122 | -5.7479
3.1681 | -34.1223 | -5.7587
3.1742 | -34.1226 | -5.7694
3.1803 | -34.1228 | -5.78
3.1864 | -34.1231 | -5.7908
3.1925 | -34.1233 | -5.8015
3.1986 | -34.1236 | -5.8124
3.2047 | -34.1239 | -5.8231
3.2108 | -34.1242 | -5.8339
3.2169 | -34.1244 | -5.8447
3.223 | -34.1247 | -5.8555
3.2292 | -34.125 | -5.8662
3.2353 | -34.1252 | -5.8772
3.2414 | -34.1255 | -5.888
3.2475 | -34.1257 | -5.8987
3.2536 | -34.126 | -5.9095
3.2597 | -34.1263 | -5.9203
3.2658 | -34.1265 | -5.9311
3.2719 | -34.1268 | -5.942
3.278 | -34.127 | -5.9528
3.2841 | -34.1273 | -5.9637
3.2902 | -34.1275 | -5.9746
3.2963 | -34.1278 | -5.9854
3.3024 | -34.1281 | -5.9963
3.3085 | -34.1283 | -6.0071
3.3146 | -34.1286 | -6.018
3.3207 | -34.1289 | -6.0289
3.3268 | -34.1291 | -6.0397
3.3329 | -34.1294 | -6.0507
3.339 | -34.1296 | -6.0616
3.3451 | -34.1299 | -6.0726
3.3512 | -34.1301 | -6.0835
3.3573 | -34.1303 | -6.0944
3.3634 | -34.1306 | -6.1053
3.3696 | -34.1309 | -6.1162
3.3757 | -34.1311 | -6.1272
3.3818 | -34.1314 | -6.138
3.3879 | -34.1316 | -6.1489
3.394 | -34.1319 | -6.1597
3.4001 | -34.1321 | -6.1705
3.4062 | -34.1324 | -6.1816
3.4123 | -34.1326 | -6.1925
3.4184 | -34.1328 | -6.2034
3.4245 | -34.1331 | -6.2143
3.4306 | -34.1333 | -6.2252
3.4367 | -34.1335 | -6.2361
3.4428 | -34.1338 | -6.247
3.4489 | -34.134 | -6.258
3.455 | -34.1343 | -6.269
3.4611 | -34.1345 | -6.2798
3.4672 | -34.1348 | -6.2907
3.4733 | -34.135 | -6.3016
3.4794 | -34.1352 | -6.3124
3.4855 | -34.1354 | -6.3233
3.4916 | -34.1357 | -6.3343
3.4977 | -34.1359 | -6.3452
3.5038 | -34.1362 | -6.3561
3.5099 | -34.1364 | -6.3671
3.5161 | -34.1366 | -6.3781
3.5222 | -34.1368 | -6.389
3.5283 | -34.1371 | -6.4
3.5344 | -34.1373 | -6.4109
3.5405 | -34.1375 | -6.4218
3.5466 | -34.1378 | -6.4327
3.5527 | -34.138 | -6.4435
3.5588 | -34.1382 | -6.4544
3.5649 | -34.1385 | -6.4653
3.571 | -34.1387 | -6.4763
3.5771 | -34.1389 | -6.4871
3.5832 | -34.1391 | -6.498
3.5893 | -34.1394 | -6.5089
3.5954 | -34.1396 | -6.5197
3.6015 | -34.1398 | -6.5305
3.6076 | -34.14 | -6.5414
3.6137 | -34.1402 | -6.5524
3.6198 | -34.1405 | -6.563
3.6259 | -34.1407 | -6.5739
3.632 | -34.1409 | -6.5847
3.6381 | -34.1411 | -6.5956
3.6442 | -34.1414 | -6.6064
3.6503 | -34.1416 | -6.6172
3.6565 | -34.1418 | -6.628
3.6626 | -34.142 | -6.6389
3.6687 | -34.1422 | -6.6497
3.6748 | -34.1424 | -6.6606
3.6809 | -34.1426 | -6.6713
3.687 | -34.1428 | -6.6821
3.6931 | -34.1431 | -6.6928
3.6992 | -34.1433 | -6.7036
3.7053 | -34.1435 | -6.7143
3.7114 | -34.1437 | -6.7251
3.7175 | -34.1439 | -6.7359
3.7236 | -34.1441 | -6.7466
3.7297 | -34.1444 | -6.7575
3.7358 | -34.1446 | -6.7683
3.7419 | -34.1448 | -6.779
3.748 | -34.145 | -6.7898
3.7541 | -34.1452 | -6.8006
3.7602 | -34.1455 | -6.8112
3.7663 | -34.1457 | -6.822
3.7724 | -34.1459 | -6.8327
3.7785 | -34.1461 | -6.8433
3.7846 | -34.1463 | -6.8541
3.7907 | -34.1466 | -6.8647
3.7969 | -34.1468 | -6.8754
3.803 | -34.147 | -6.8861
3.8091 | -34.1472 | -6.897
3.8152 | -34.1474 | -6.9075
3.8213 | -34.1477 | -6.9182
3.8274 | -34.1479 | -6.929
3.8335 | -34.1481 | -6.9395
3.8396 | -34.1484 | -6.9502
3.8457 | -34.1486 | -6.9608
3.8518 | -34.1488 | -6.9714
3.8579 | -34.149 | -6.982
3.864 | -34.1493 | -6.9927
3.8701 | -34.1495 | -7.0033
3.8762 | -34.1497 | -7.014
3.8823 | -34.1499 | -7.0247
3.8884 | -34.1502 | -7.0354
3.8945 | -34.1504 | -7.046
3.9006 | -34.1507 | -7.0567
3.9067 | -34.1509 | -7.0673
3.9128 | -34.1511 | -7.078
3.9189 | -34.1514 | -7.0886
3.925 | -34.1516 | -7.0993
3.9311 | -34.1518 | -7.11
3.9372 | -34.1521 | -7.1206
3.9434 | -34.1523 | -7.1313
3.9495 | -34.1525 | -7.1419
3.9556 | -34.1528 | -7.1525
3.9617 | -34.153 | -7.1633
3.9678 | -34.1532 | -7.1739
3.9739 | -34.1535 | -7.1846
3.98 | -34.1537 | -7.1953
3.9861 | -34.154 | -7.2059
3.9922 | -34.1542 | -7.2166
3.9983 | -34.1545 | -7.2272
4.0044 | -34.1547 | -7.238
4.0105 | -34.155 | -7.2486
4.0166 | -34.1552 | -7.2592
4.0227 | -34.1554 | -7.2699
4.0288 | -34.1557 | -7.2806
4.0349 | -34.1559 | -7.2913
4.041 | -34.1562 | -7.3019
4.0471 | -34.1564 | -7.3126
4.0532 | -34.1567 | -7.3232
4.0593 | -34.1569 | -7.3338
4.0654 | -34.1572 | -7.3444
4.0715 | -34.1574 | -7.3551
4.0776 | -34.1577 | -7.3656
4.0838 | -34.158 | -7.3762
4.0899 | -34.1582 | -7.3869
4.096 | -34.1584 | -7.3974
4.1021 | -34.1587 | -7.4081
4.1082 | -34.1589 | -7.4188
4.1143 | -34.1592 | -7.4295
4.1204 | -34.1594 | -7.4402
4.1265 | -34.1597 | -7.4508
4.1326 | -34.1599 | -7.4614
4.1387 | -34.1602 | -7.4719
4.1448 | -34.1605 | -7.4826
4.1509 | -34.1608 | -7.4932
4.157 | -34.161 | -7.5039
4.1631 | -34.1613 | -7.5145
4.1692 | -34.1615 | -7.5251
4.1753 | -34.1618 | -7.5356
4.1814 | -34.1621 | -7.5463
4.1875 | -34.1623 | -7.557
4.1936 | -34.1626 | -7.5677
4.1997 | -34.1628 | -7.5784
4.2058 | -34.1631 | -7.589
4.2119 | -34.1634 | -7.5997
4.218 | -34.1636 | -7.6104
4.2241 | -34.1639 | -7.6211
4.2303 | -34.1641 | -7.6317
4.2364 | -34.1644 | -7.6424
4.2425 | -34.1647 | -7.6531
4.2486 | -34.165 | -7.6637
4.2547 | -34.1652 | -7.6744
4.2608 | -34.1655 | -7.6851
4.2669 | -34.1658 | -7.6957
4.273 | -34.166 | -7.7063
4.2791 | -34.1663 | -7.717
4.2852 | -34.1666 | -7.7276
4.2913 | -34.1668 | -7.7383
4.2974 | -34.1671 | -7.749
4.3035 | -34.1674 | -7.7596
4.3096 | -34.1677 | -7.7703
4.3157 | -34.1679 | -7.781
4.3218 | -34.1682 | -7.7917
4.3279 | -34.1685 | -7.8023
4.334 | -34.1687 | -7.8131
4.3401 | -34.169 | -7.8237
4.3462 | -34.1693 | -7.8343
4.3523 | -34.1695 | -7.8449
4.3584 | -34.1698 | -7.8556
4.3645 | -34.17 | -7.8662
4.3707 | -34.1703 | -7.8769
4.3768 | -34.1706 | -7.8876
4.3829 | -34.1709 | -7.8983
4.389 | -34.1711 | -7.909
4.3951 | -34.1714 | -7.9197
4.4012 | -34.1717 | -7.9304
4.4073 | -34.1719 | -7.9412
4.4134 | -34.1722 | -7.9518
4.4195 | -34.1725 | -7.9624
4.4256 | -34.1727 | -7.973
4.4317 | -34.173 | -7.9837
4.4378 | -34.1733 | -7.9944
4.4439 | -34.1735 | -8.0051
4.45 | -34.1738 | -8.0157
4.4561 | -34.174 | -8.0265
4.4622 | -34.1743 | -8.0372
4.4683 | -34.1746 | -8.0479
4.4744 | -34.1749 | -8.0585
4.4805 | -34.1751 | -8.0691
4.4866 | -34.1754 | -8.0798
4.4927 | -34.1756 | -8.0904
4.4988 | -34.1759 | -8.101
4.5049 | -34.1761 | -8.1116
4.511 | -34.1764 | -8.1223
4.5172 | -34.1767 | -8.1328
4.5233 | -34.1769 | -8.1435
4.5294 | -34.1772 | -8.1541
4.5355 | -34.1774 | -8.1648
4.5416 | -34.1777 | -8.1754
4.5477 | -34.1779 | -8.1861
4.5538 | -34.1782 | -8.1967
4.5599 | -34.1784 | -8.2073
4.566 | -34.1787 | -8.2179
4.5721 | -34.1789 | -8.2285
4.5782 | -34.1792 | -8.239
4.5843 | -34.1795 | -8.2496
4.5904 | -34.1797 | -8.2602
4.5965 | -34.18 | -8.2709
4.6026 | -34.1802 | -8.2815
4.6087 | -34.1805 | -8.2921
4.6148 | -34.1807 | -8.3027
4.6209 | -34.181 | -8.3133
4.627 | -34.1813 | -8.3239
4.6331 | -34.1815 | -8.3346
4.6392 | -34.1818 | -8.3452
4.6453 | -34.182 | -8.3557
4.6514 | -34.1823 | -8.3662
4.6576 | -34.1825 | -8.3769
4.6637 | -34.1828 | -8.3874
4.6698 | -34.183 | -8.398
4.6759 | -34.1833 | -8.4085
4.682 | -34.1835 | -8.419
4.6881 | -34.1838 | -8.4296
4.6942 | -34.184 | -8.4402
4.7003 | -34.1843 | -8.4508
4.7064 | -34.1845 | -8.4612
4.7125 | -34.1848 | -8.4717
4.7186 | -34.1851 | -8.4822
4.7247 | -34.1853 | -8.4927
4.7308 | -34.1856 | -8.5032
4.7369 | -34.1858 | -8.5137
4.743 | -34.1861 | -8.5241
4.7491 | -34.1863 | -8.5346
4.7552 | -34.1866 | -8.5451
4.7613 | -34.1868 | -8.5557
4.7674 | -34.1871 | -8.5662
4.7735 | -34.1874 | -8.5768
4.7796 | -34.1876 | -8.587
4.7857 | -34.1879 | -8.5975
4.7918 | -34.1881 | -8.608
4.7979 | -34.1883 | -8.6184
4.8041 | -34.1886 | -8.6289
4.8102 | -34.1888 | -8.6394
4.8163 | -34.1891 | -8.6498
4.8224 | -34.1893 | -8.6603
4.8285 | -34.1896 | -8.6708
4.8346 | -34.1898 | -8.6813
4.8407 | -34.1901 | -8.6917
4.8468 | -34.1903 | -8.7021
4.8529 | -34.1906 | -8.7125
4.859 | -34.1908 | -8.7229
4.8651 | -34.1911 | -8.7333
4.8712 | -34.1913 | -8.7438
4.8773 | -34.1916 | -8.7543
4.8834 | -34.1918 | -8.7647
4.8895 | -34.1921 | -8.7751
4.8956 | -34.1923 | -8.7855
4.9017 | -34.1926 | -8.7959
4.9078 | -34.1928 | -8.8063
4.9139 | -34.193 | -8.8166
4.92 | -34.1933 | -8.8269
4.9261 | -34.1935 | -8.8373
4.9322 | -34.1938 | -8.8477
4.9383 | -34.194 | -8.8581
4.9445 | -34.1943 | -8.8684
4.9506 | -34.1945 | -8.8788
4.9567 | -34.1948 | -8.8891
4.9628 | -34.195 | -8.8994
4.9689 | -34.1953 | -8.9098
4.975 | -34.1955 | -8.9202
4.9811 | -34.1958 | -8.9306
4.9872 | -34.196 | -8.9409
4.9933 | -34.1962 | -8.9511
4.9994 | -34.1965 | -8.9614
5.0055 | -34.1967 | -8.9717
5.0116 | -34.197 | -8.982
5.0177 | -34.1972 | -8.9923
5.0238 | -34.1975 | -9.0025
5.0299 | -34.1977 | -9.0128
5.036 | -34.198 | -9.0232
5.0421 | -34.1982 | -9.0335
5.0482 | -34.1985 | -9.0439
5.0543 | -34.1987 | -9.0542
5.0604 | -34.1989 | -9.0645
5.0665 | -34.1992 | -9.0748
5.0726 | -34.1994 | -9.0851
5.0787 | -34.1997 | -9.0954
5.0848 | -34.1999 | -9.1056
5.091 | -34.2002 | -9.116
5.0971 | -34.2004 | -9.1261
5.1032 | -34.2007 | -9.1365
5.1093 | -34.2009 | -9.1467
5.1154 | -34.2012 | -9.1571
5.1215 | -34.2014 | -9.1674
5.1276 | -34.2017 | -9.1778
5.1337 | -34.2019 | -9.1879
5.1398 | -34.2022 | -9.1982
5.1459 | -34.2024 | -9.2084
5.152 | -34.2027 | -9.2187
5.1581 | -34.2029 | -9.229
5.1642 | -34.2032 | -9.2393
5.1703 | -34.2034 | -9.2496
5.1764 | -34.2037 | -9.2599
5.1825 | -34.2039 | -9.2701
5.1886 | -34.2042 | -9.2805
5.1947 | -34.2044 | -9.2908
5.2008 | -34.2047 | -9.3011
5.2069 | -34.2049 | -9.3114
5.213 | -34.2052 | -9.3218
5.2191 | -34.2054 | -9.332
5.2252 | -34.2057 | -9.3423
5.2314 | -34.2059 | -9.3525
5.2375 | -34.2062 | -9.3628
5.2436 | -34.2065 | -9.3733
5.2497 | -34.2067 | -9.3836
5.2558 | -34.207 | -9.3938
5.2619 | -34.2072 | -9.4042
5.268 | -34.2075 | -9.4144
5.2741 | -34.2078 | -9.4246
5.2802 | -34.208 | -9.4349
5.2863 | -34.2083 | -9.4451
5.2924 | -34.2085 | -9.4554
5.2985 | -34.2088 | -9.4657
5.3046 | -34.209 | -9.476
5.3107 | -34.2093 | -9.4863
5.3168 | -34.2096 | -9.4967
5.3229 | -34.2098 | -9.5069
5.329 | -34.2101 | -9.5172
5.3351 | -34.2104 | -9.5275
5.3412 | -34.2106 | -9.5379
5.3473 | -34.2109 | -9.5482
5.3534 | -34.2112 | -9.5585
5.3595 | -34.2114 | -9.5689
5.3656 | -34.2117 | -9.5791
5.3717 | -34.2119 | -9.5896
5.3779 | -34.2122 | -9.5999
5.384 | -34.2125 | -9.6102
5.3901 | -34.2128 | -9.6205
5.3962 | -34.213 | -9.6309
5.4023 | -34.2133 | -9.6411
5.4084 | -34.2135 | -9.6514
5.4145 | -34.2138 | -9.6618
5.4206 | -34.2141 | -9.6721
5.4267 | -34.2144 | -9.6823
5.4328 | -34.2146 | -9.6927
5.4389 | -34.2149 | -9.7031
5.445 | -34.2152 | -9.7135
5.4511 | -34.2154 | -9.7237
5.4572 | -34.2157 | -9.734
5.4633 | -34.216 | -9.7443
5.4694 | -34.2163 | -9.7547
5.4755 | -34.2165 | -9.765
5.4816 | -34.2168 | -9.7753
5.4877 | -34.2171 | -9.7857
5.4938 | -34.2174 | -9.7961
5.4999 | -34.2176 | -9.8065
5.506 | -34.2179 | -9.8168
5.5121 | -34.2182 | -9.8271
5.5183 | -34.2185 | -9.8377
5.5244 | -34.2187 | -9.8481
5.5305 | -34.219 | -9.8584
5.5366 | -34.2193 | -9.8688
5.5427 | -34.2196 | -9.8792
5.5488 | -34.2198 | -9.8896
5.5549 | -34.2201 | -9.9002
5.561 | -34.2204 | -9.9104
5.5671 | -34.2207 | -9.9207
5.5732 | -34.221 | -9.9312
5.5793 | -34.2213 | -9.9416
5.5854 | -34.2215 | -9.9519
5.5915 | -34.2219 | -9.9623
5.5976 | -34.2221 | -9.9727
5.6037 | -34.2224 | -9.9831
5.6098 | -34.2227 | -9.9935
5.6159 | -34.223 | -10.0039
5.622 | -34.2232 | -10.0143
5.6281 | -34.2235 | -10.0248
5.6342 | -34.2238 | -10.0352
5.6403 | -34.2241 | -10.0456
5.6464 | -34.2244 | -10.056
5.6525 | -34.2247 | -10.0664
5.6586 | -34.225 | -10.0769
5.6648 | -34.2253 | -10.0872
5.6709 | -34.2256 | -10.0975
5.677 | -34.2259 | -10.1079
5.6831 | -34.2262 | -10.1184
5.6892 | -34.2265 | -10.1289
5.6953 | -34.2268 | -10.1394
5.7014 | -34.2271 | -10.1499
5.7075 | -34.2274 | -10.1603
5.7136 | -34.2276 | -10.1708
5.7197 | -34.2279 | -10.1812
5.7258 | -34.2282 | -10.1916
5.7319 | -34.2285 | -10.202
5.738 | -34.2289 | -10.2124
5.7441 | -34.2292 | -10.2229
5.7502 | -34.2294 | -10.2332
5.7563 | -34.2297 | -10.2437
5.7624 | -34.23 | -10.2541
5.7685 | -34.2303 | -10.2645
5.7746 | -34.2306 | -10.275
5.7807 | -34.2309 | -10.2855
5.7868 | -34.2312 | -10.2959
5.7929 | -34.2316 | -10.3064
5.799 | -34.2319 | -10.3169
5.8052 | -34.2321 | -10.3273
5.8113 | -34.2325 | -10.3378
5.8174 | -34.2328 | -10.3483
5.8235 | -34.2331 | -10.3587
5.8296 | -34.2334 | -10.3693
5.8357 | -34.2337 | -10.3798
5.8418 | -34.234 | -10.3902
5.8479 | -34.2343 | -10.4007
5.854 | -34.2346 | -10.4112
5.8601 | -34.2349 | -10.4216
5.8662 | -34.2352 | -10.4319
5.8723 | -34.2355 | -10.4424
5.8784 | -34.2358 | -10.453
5.8845 | -34.2361 | -10.4634
5.8906 | -34.2364 | -10.4738
5.8967 | -34.2367 | -10.4842
5.9028 | -34.2369 | -10.4947
5.9089 | -34.2373 | -10.505
5.915 | -34.2375 | -10.5154
5.9211 | -34.2378 | -10.5259
5.9272 | -34.2381 | -10.5364
5.9333 | -34.2384 | -10.5468
5.9394 | -34.2387 | -10.5573
5.9455 | -34.239 | -10.5677
5.9517 | -34.2393 | -10.578
5.9578 | -34.2396 | -10.5885
5.9639 | -34.2399 | -10.5989
5.97 | -34.2402 | -10.6092
5.9761 | -34.2405 | -10.6197
5.9822 | -34.2408 | -10.63
5.9883 | -34.2411 | -10.6404
5.9944 | -34.2414 | -10.6508
6.0005 | -34.2417 | -10.6612
6.0066 | -34.242 | -10.6715
6.0127 | -34.2424 | -10.6818
6.0188 | -34.2427 | -10.6923
6.0249 | -34.243 | -10.7027
6.031 | -34.2433 | -10.7132
6.0371 | -34.2436 | -10.7236
6.0432 | -34.2439 | -10.7341
6.0493 | -34.2442 | -10.7444
6.0554 | -34.2445 | -10.7548
6.0615 | -34.2448 | -10.7652
6.0676 | -34.2451 | -10.7756
6.0737 | -34.2454 | -10.786
6.0798 | -34.2457 | -10.7965
6.0859 | -34.246 | -10.8067
6.0921 | -34.2463 | -10.8171
6.0982 | -34.2466 | -10.8275
6.1043 | -34.2469 | -10.8378
6.1104 | -34.2473 | -10.8481
6.1165 | -34.2476 | -10.8584
6.1226 | -34.2479 | -10.8688
6.1287 | -34.2482 | -10.8792
6.1348 | -34.2485 | -10.8896
6.1409 | -34.2488 | -10.8999
6.147 | -34.2491 | -10.9102
6.1531 | -34.2494 | -10.9205
6.1592 | -34.2497 | -10.9309
6.1653 | -34.25 | -10.9412
6.1714 | -34.2503 | -10.9515
6.1775 | -34.2507 | -10.9618
6.1836 | -34.251 | -10.9723
6.1897 | -34.2513 | -10.9824
6.1958 | -34.2516 | -10.9928
6.2019 | -34.2519 | -11.0031
6.208 | -34.2522 | -11.0134
6.2141 | -34.2525 | -11.0237
6.2202 | -34.2528 | -11.034
6.2263 | -34.2532 | -11.0441
6.2325 | -34.2535 | -11.0545
6.2386 | -34.2538 | -11.0647
6.2447 | -34.2541 | -11.075
6.2508 | -34.2544 | -11.0852
6.2569 | -34.2547 | -11.0954
6.263 | -34.255 | -11.1057
6.2691 | -34.2554 | -11.116
6.2752 | -34.2557 | -11.1263
6.2813 | -34.256 | -11.1366
6.2874 | -34.2563 | -11.1469
6.2935 | -34.2566 | -11.157
6.2996 | -34.2569 | -11.1674
6.3057 | -34.2572 | -11.1776
6.3118 | -34.2575 | -11.1879
6.3179 | -34.2579 | -11.1982
6.324 | -34.2582 | -11.2084
6.3301 | -34.2585 | -11.2184
6.3362 | -34.2588 | -11.2287
6.3423 | -34.2591 | -11.2389
6.3484 | -34.2594 | -11.2491
6.3545 | -34.2597 | -11.2593
6.3606 | -34.26 | -11.2696
6.3667 | -34.2603 | -11.2798
6.3728 | -34.2607 | -11.2901
6.379 | -34.261 | -11.3003
6.3851 | -34.2613 | -11.3105
6.3912 | -34.2616 | -11.3208
6.3973 | -34.2619 | -11.3309
6.4034 | -34.2622 | -11.3411
6.4095 | -34.2626 | -11.3513
6.4156 | -34.2629 | -11.3614
6.4217 | -34.2632 | -11.3716
6.4278 | -34.2635 | -11.3819
6.4339 | -34.2638 | -11.3919
6.44 | -34.2641 | -11.4021
6.4461 | -34.2645 | -11.4124
6.4522 | -34.2648 | -11.4226
6.4583 | -34.2651 | -11.4328
6.4644 | -34.2654 | -11.443
6.4705 | -34.2657 | -11.453
6.4766 | -34.2661 | -11.4631
6.4827 | -34.2664 | -11.4733
6.4888 | -34.2667 | -11.4836
6.4949 | -34.267 | -11.4937
6.501 | -34.2673 | -11.5038
6.5071 | -34.2676 | -11.5138
6.5132 | -34.268 | -11.524
6.5194 | -34.2683 | -11.5342
6.5255 | -34.2686 | -11.5443
6.5316 | -34.2689 | -11.5545
6.5377 | -34.2692 | -11.5646
6.5438 | -34.2696 | -11.5745
6.5499 | -34.2699 | -11.5846
6.556 | -34.2702 | -11.5948
6.5621 | -34.2705 | -11.6049
6.5682 | -34.2708 | -11.615
6.5743 | -34.2712 | -11.625
6.5804 | -34.2715 | -11.6351
6.5865 | -34.2718 | -11.6451
6.5926 | -34.2721 | -11.6552
6.5987 | -34.2725 | -11.6652
6.6048 | -34.2728 | -11.6753
6.6109 | -34.2731 | -11.6852
6.617 | -34.2734 | -11.6953
6.6231 | -34.2737 | -11.7054
6.6292 | -34.274 | -11.7154
6.6353 | -34.2744 | -11.7256
6.6414 | -34.2747 | -11.7356
6.6475 | -34.275 | -11.7456
6.6536 | -34.2753 | -11.7556
6.6597 | -34.2756 | -11.7657
6.6659 | -34.2759 | -11.7758
6.672 | -34.2763 | -11.7858
6.6781 | -34.2766 | -11.7959
6.6842 | -34.2769 | -11.8058
6.6903 | -34.2772 | -11.8159
6.6964 | -34.2776 | -11.8259
6.7025 | -34.2779 | -11.8359
6.7086 | -34.2782 | -11.8458
6.7147 | -34.2785 | -11.8559
6.7208 | -34.2788 | -11.8659
6.7269 | -34.2792 | -11.876
6.733 | -34.2795 | -11.886
6.7391 | -34.2798 | -11.8959
6.7452 | -34.2801 | -11.9059
6.7513 | -34.2804 | -11.9158
6.7574 | -34.2807 | -11.9258
6.7635 | -34.2811 | -11.9358
6.7696 | -34.2814 | -11.9457
6.7757 | -34.2817 | -11.9557
6.7818 | -34.282 | -11.9657
6.7879 | -34.2824 | -11.9755
6.794 | -34.2827 | -11.9855
6.8001 | -34.283 | -11.9955
6.8063 | -34.2833 | -12.0055
6.8124 | -34.2836 | -12.0155
6.8185 | -34.284 | -12.0254
6.8246 | -34.2843 | -12.0353
6.8307 | -34.2846 | -12.0453
6.8368 | -34.285 | -12.0552
6.8429 | -34.2853 | -12.0651
6.849 | -34.2856 | -12.0751
6.8551 | -34.286 | -12.0848
6.8612 | -34.2863 | -12.0948
6.8673 | -34.2866 | -12.1047
6.8734 | -34.2869 | -12.1146
6.8795 | -34.2872 | -12.1246
6.8856 | -34.2876 | -12.1345
6.8917 | -34.2879 | -12.1444
6.8978 | -34.2882 | -12.1543
6.9039 | -34.2885 | -12.1642
6.91 | -34.2888 | -12.1741
6.9161 | -34.2891 | -12.184
6.9222 | -34.2895 | -12.194
6.9283 | -34.2898 | -12.2039
6.9344 | -34.2901 | -12.2138
6.9405 | -34.2905 | -12.2237
6.9466 | -34.2908 | -12.2336
6.9528 | -34.2911 | -12.2435
6.9589 | -34.2915 | -12.2533
6.965 | -34.2918 | -12.2632
6.9711 | -34.2921 | -12.273
6.9772 | -34.2925 | -12.283
6.9833 | -34.2928 | -12.2929
6.9894 | -34.2931 | -12.3028
6.9955 | -34.2934 | -12.3126
7.0016 | -34.2937 | -12.3225
7.0077 | -34.2941 | -12.3325
7.0138 | -34.2944 | -12.3423
7.0199 | -34.2947 | -12.3522
7.026 | -34.295 | -12.362
7.0321 | -34.2954 | -12.3719
7.0382 | -34.2957 | -12.3818
7.0443 | -34.296 | -12.3917
7.0504 | -34.2963 | -12.4017
7.0565 | -34.2967 | -12.4116
7.0626 | -34.297 | -12.4215
7.0687 | -34.2973 | -12.4314
7.0748 | -34.2976 | -12.4412
7.0809 | -34.298 | -12.4511
7.087 | -34.2983 | -12.4609
7.0932 | -34.2986 | -12.4709
7.0993 | -34.299 | -12.4807
7.1054 | -34.2993 | -12.4905
7.1115 | -34.2996 | -12.5003
7.1176 | -34.2999 | -12.5103
7.1237 | -34.3002 | -12.5203
7.1298 | -34.3006 | -12.5301
7.1359 | -34.3009 | -12.54
7.142 | -34.3012 | -12.5499
7.1481 | -34.3016 | -12.5599
7.1542 | -34.3019 | -12.5698
7.1603 | -34.3022 | -12.5796
7.1664 | -34.3026 | -12.5895
7.1725 | -34.3029 | -12.5994
7.1786 | -34.3032 | -12.6092
7.1847 | -34.3035 | -12.6191
7.1908 | -34.3039 | -12.629
7.1969 | -34.3042 | -12.639
7.203 | -34.3045 | -12.6488
7.2091 | -34.3048 | -12.6587
7.2152 | -34.3052 | -12.6687
7.2213 | -34.3055 | -12.6786
7.2274 | -34.3058 | -12.6884
7.2335 | -34.3061 | -12.6984
7.2397 | -34.3065 | -12.7083
7.2458 | -34.3068 | -12.7181
7.2519 | -34.3071 | -12.728
7.258 | -34.3075 | -12.7379
7.2641 | -34.3078 | -12.7477
7.2702 | -34.3081 | -12.7576
7.2763 | -34.3084 | -12.7673
7.2824 | -34.3088 | -12.7772
7.2885 | -34.3091 | -12.7871
7.2946 | -34.3094 | -12.797
7.3007 | -34.3098 | -12.8069
7.3068 | -34.3101 | -12.8169
7.3129 | -34.3104 | -12.8268
7.319 | -34.3108 | -12.8367
7.3251 | -34.3111 | -12.8466
7.3312 | -34.3114 | -12.8565
7.3373 | -34.3118 | -12.8664
7.3434 | -34.3121 | -12.8762
7.3495 | -34.3124 | -12.8862
7.3556 | -34.3128 | -12.8961
7.3617 | -34.3131 | -12.9059
7.3678 | -34.3134 | -12.9158
7.3739 | -34.3137 | -12.9258
7.3801 | -34.3141 | -12.9357
7.3862 | -34.3144 | -12.9456
7.3923 | -34.3148 | -12.9554
7.3984 | -34.3151 | -12.9653
7.4045 | -34.3155 | -12.9752
7.4106 | -34.3158 | -12.9851
7.4167 | -34.3161 | -12.995
7.4228 | -34.3165 | -13.0048
7.4289 | -34.3168 | -13.0147
7.435 | -34.3172 | -13.0247
7.4411 | -34.3175 | -13.0347
7.4472 | -34.3178 | -13.0446
7.4533 | -34.3182 | -13.0545
7.4594 | -34.3185 | -13.0644
7.4655 | -34.3188 | -13.0743
7.4716 | -34.3192 | -13.0841
7.4777 | -34.3195 | -13.094
7.4838 | -34.3198 | -13.1039
7.4899 | -34.3202 | -13.1139
7.496 | -34.3205 | -13.124
7.5021 | -34.3209 | -13.1339
7.5082 | -34.3212 | -13.1437
7.5143 | -34.3215 | -13.1537
7.5204 | -34.3219 | -13.1636
7.5266 | -34.3222 | -13.1737
7.5327 | -34.3226 | -13.1835
7.5388 | -34.3229 | -13.1936
7.5449 | -34.3233 | -13.2033
7.551 | -34.3236 | -13.2133
7.5571 | -34.324 | -13.2232
7.5632 | -34.3243 | -13.2331
7.5693 | -34.3247 | -13.243
7.5754 | -34.325 | -13.253
7.5815 | -34.3253 | -13.2629
7.5876 | -34.3257 | -13.2728
7.5937 | -34.326 | -13.2827
7.5998 | -34.3264 | -13.2925
7.6059 | -34.3267 | -13.3024
7.612 | -34.3271 | -13.3123
7.6181 | -34.3274 | -13.3222
7.6242 | -34.3278 | -13.3321
7.6303 | -34.3281 | -13.3421
7.6364 | -34.3285 | -13.3521
7.6425 | -34.3288 | -13.3619
7.6486 | -34.3292 | -13.3719
7.6547 | -34.3295 | -13.3817
7.6608 | -34.3299 | -13.3916
7.667 | -34.3302 | -13.4016
7.6731 | -34.3306 | -13.4115
7.6792 | -34.331 | -13.4214
7.6853 | -34.3313 | -13.4313
7.6914 | -34.3317 | -13.4412
7.6975 | -34.332 | -13.4512
7.7036 | -34.3324 | -13.4611
7.7097 | -34.3328 | -13.4709
7.7158 | -34.3331 | -13.4808
7.7219 | -34.3335 | -13.4908
7.728 | -34.3338 | -13.5007
7.7341 | -34.3342 | -13.5107
7.7402 | -34.3345 | -13.5206
7.7463 | -34.3349 | -13.5305
7.7524 | -34.3353 | -13.5404
7.7585 | -34.3356 | -13.5504
7.7646 | -34.336 | -13.5603
7.7707 | -34.3364 | -13.5703
7.7768 | -34.3367 | -13.5803
7.7829 | -34.3371 | -13.5902
7.789 | -34.3375 | -13.6002
7.7951 | -34.3378 | -13.6101
7.8012 | -34.3382 | -13.6199
7.8073 | -34.3386 | -13.6297
7.8135 | -34.3389 | -13.6397
7.8196 | -34.3393 | -13.6496
7.8257 | -34.3397 | -13.6596
7.8318 | -34.34 | -13.6696
7.8379 | -34.3404 | -13.6795
7.844 | -34.3407 | -13.6895
7.8501 | -34.3411 | -13.6994
7.8562 | -34.3415 | -13.7093
7.8623 | -34.3419 | -13.7191
7.8684 | -34.3422 | -13.729
7.8745 | -34.3426 | -13.739
7.8806 | -34.343 | -13.7489
7.8867 | -34.3434 | -13.7588
7.8928 | -34.3437 | -13.7686
7.8989 | -34.3441 | -13.7785
7.905 | -34.3445 | -13.7884
7.9111 | -34.3449 | -13.7983
7.9172 | -34.3452 | -13.8082
7.9233 | -34.3456 | -13.8182
7.9294 | -34.346 | -13.8281
7.9355 | -34.3463 | -13.8381
7.9416 | -34.3467 | -13.848
7.9477 | -34.3471 | -13.8579
7.9539 | -34.3475 | -13.8678
7.96 | -34.3478 | -13.8779
7.9661 | -34.3482 | -13.8878
7.9722 | -34.3486 | -13.8977
7.9783 | -34.349 | -13.9076
7.9844 | -34.3494 | -13.9175
7.9905 | -34.3497 | -13.9274
7.9966 | -34.3501 | -13.9374
8.0027 | -34.3505 | -13.9474
8.0088 | -34.3509 | -13.9573
8.0149 | -34.3512 | -13.9672
8.021 | -34.3516 | -13.9771
8.0271 | -34.352 | -13.987
8.0332 | -34.3524 | -13.997
8.0393 | -34.3528 | -14.0068
8.0454 | -34.3532 | -14.0167
8.0515 | -34.3536 | -14.0264
8.0576 | -34.3539 | -14.0363
8.0637 | -34.3543 | -14.0464
8.0698 | -34.3547 | -14.0562
8.0759 | -34.355 | -14.0662
8.082 | -34.3554 | -14.076
8.0881 | -34.3558 | -14.086
8.0942 | -34.3562 | -14.0957
8.1004 | -34.3566 | -14.1056
8.1065 | -34.357 | -14.1156
8.1126 | -34.3573 | -14.1255
8.1187 | -34.3577 | -14.1354
8.1248 | -34.3581 | -14.1453
8.1309 | -34.3585 | -14.1551
8.137 | -34.3589 | -14.1651
8.1431 | -34.3592 | -14.1751
8.1492 | -34.3596 | -14.185
8.1553 | -34.36 | -14.1948
8.1614 | -34.3604 | -14.2047
8.1675 | -34.3608 | -14.2146
8.1736 | -34.3612 | -14.2245
8.1797 | -34.3615 | -14.2344
8.1858 | -34.3619 | -14.2443
8.1919 | -34.3623 | -14.2542
8.198 | -34.3627 | -14.2641
8.2041 | -34.3631 | -14.2739
8.2102 | -34.3635 | -14.2838
8.2163 | -34.3639 | -14.2937
8.2224 | -34.3642 | -14.3035
8.2285 | -34.3646 | -14.3134
8.2346 | -34.365 | -14.3234
8.2408 | -34.3654 | -14.3332
8.2469 | -34.3658 | -14.3431
8.253 | -34.3662 | -14.3529
8.2591 | -34.3666 | -14.3627
8.2652 | -34.367 | -14.3725
8.2713 | -34.3674 | -14.3823
8.2774 | -34.3678 | -14.3921
8.2835 | -34.3682 | -14.402
8.2896 | -34.3685 | -14.4119
8.2957 | -34.3689 | -14.4218
8.3018 | -34.3693 | -14.4316
8.3079 | -34.3697 | -14.4414
8.314 | -34.3701 | -14.4512
8.3201 | -34.3705 | -14.461
8.3262 | -34.3709 | -14.4709
8.3323 | -34.3713 | -14.4806
8.3384 | -34.3717 | -14.4905
8.3445 | -34.3721 | -14.5005
8.3506 | -34.3725 | -14.5103
8.3567 | -34.3729 | -14.5201
8.3628 | -34.3733 | -14.53
8.3689 | -34.3737 | -14.5397
8.375 | -34.3741 | -14.5496
8.3812 | -34.3745 | -14.5595
8.3873 | -34.3748 | -14.5694
8.3934 | -34.3752 | -14.5792
8.3995 | -34.3756 | -14.5891
8.4056 | -34.3761 | -14.5989
8.4117 | -34.3764 | -14.6087
8.4178 | -34.3768 | -14.6187
8.4239 | -34.3772 | -14.6285
8.43 | -34.3776 | -14.6383
8.4361 | -34.378 | -14.6483
8.4422 | -34.3784 | -14.6579
8.4483 | -34.3788 | -14.6677
8.4544 | -34.3792 | -14.6775
8.4605 | -34.3796 | -14.6874
8.4666 | -34.38 | -14.6972
8.4727 | -34.3804 | -14.707
8.4788 | -34.3808 | -14.7168
8.4849 | -34.3812 | -14.7266
8.491 | -34.3816 | -14.7365
8.4971 | -34.382 | -14.7462
8.5032 | -34.3823 | -14.756
8.5093 | -34.3827 | -14.7656
8.5154 | -34.3831 | -14.7754
8.5215 | -34.3835 | -14.7853
8.5277 | -34.3839 | -14.7951
8.5338 | -34.3843 | -14.8048
8.5399 | -34.3847 | -14.8147
8.546 | -34.3851 | -14.8242
8.5521 | -34.3855 | -14.8341
8.5582 | -34.3858 | -14.8439
8.5643 | -34.3862 | -14.8537
8.5704 | -34.3866 | -14.8634
8.5765 | -34.387 | -14.8732
8.5826 | -34.3874 | -14.8829
8.5887 | -34.3878 | -14.8926
8.5948 | -34.3882 | -14.9024
8.6009 | -34.3886 | -14.9121
8.607 | -34.389 | -14.9219
8.6131 | -34.3894 | -14.9316
8.6192 | -34.3898 | -14.9412
8.6253 | -34.3902 | -14.951
8.6314 | -34.3906 | -14.9608
8.6375 | -34.3909 | -14.9707
8.6436 | -34.3913 | -14.9803
8.6497 | -34.3917 | -14.9899
8.6558 | -34.3921 | -14.9997
8.6619 | -34.3925 | -15.0094
8.6681 | -34.3929 | -15.0191
8.6742 | -34.3933 | -15.0288
8.6803 | -34.3937 | -15.0387
8.6864 | -34.3941 | -15.0482
8.6925 | -34.3944 | -15.058
8.6986 | -34.3948 | -15.0677
8.7047 | -34.3952 | -15.0776
8.7108 | -34.3956 | -15.0873
8.7169 | -34.396 | -15.0969
8.723 | -34.3964 | -15.1065
8.7291 | -34.3968 | -15.1162
8.7352 | -34.3972 | -15.1259
8.7413 | -34.3976 | -15.1357
8.7474 | -34.3979 | -15.1454
8.7535 | -34.3983 | -15.1551
8.7596 | -34.3988 | -15.1647
8.7657 | -34.3992 | -15.1743
8.7718 | -34.3996 | -15.1841
8.7779 | -34.4 | -15.1938
8.784 | -34.4004 | -15.2035
8.7901 | -34.4008 | -15.2129
8.7962 | -34.4011 | -15.2226
8.8023 | -34.4015 | -15.2323
8.8084 | -34.4019 | -15.242
8.8146 | -34.4023 | -15.2517
8.8207 | -34.4027 | -15.2614
8.8268 | -34.4031 | -15.2708
8.8329 | -34.4035 | -15.2805
8.839 | -34.4039 | -15.2901
8.8451 | -34.4043 | -15.2999
8.8512 | -34.4046 | -15.3096
8.8573 | -34.405 | -15.3193
8.8634 | -34.4054 | -15.3287
8.8695 | -34.4058 | -15.3384
8.8756 | -34.4062 | -15.3481
8.8817 | -34.4065 | -15.3577
8.8878 | -34.4069 | -15.3673
8.8939 | -34.4073 | -15.377
8.9 | -34.4077 | -15.3865
8.9061 | -34.4081 | -15.3962
8.9122 | -34.4085 | -15.4058
8.9183 | -34.4089 | -15.4155
8.9244 | -34.4093 | -15.425
8.9305 | -34.4096 | -15.4344
8.9366 | -34.41 | -15.444
8.9427 | -34.4104 | -15.4536
8.9488 | -34.4108 | -15.4632
8.955 | -34.4112 | -15.4729
8.9611 | -34.4116 | -15.4824
8.9672 | -34.412 | -15.4919
8.9733 | -34.4124 | -15.5015
8.9794 | -34.4128 | -15.511
8.9855 | -34.4131 | -15.5206
8.9916 | -34.4135 | -15.5303
8.9977 | -34.4139 | -15.5399
9.0038 | -34.4143 | -15.5494
9.0099 | -34.4147 | -15.5589
9.016 | -34.4151 | -15.5686
9.0221 | -34.4155 | -15.5782
9.0282 | -34.4159 | -15.5879
9.0343 | -34.4162 | -15.5974
9.0404 | -34.4167 | -15.6069
9.0465 | -34.4171 | -15.6165
9.0526 | -34.4174 | -15.626
9.0587 | -34.4178 | -15.6355
9.0648 | -34.4182 | -15.6451
9.0709 | -34.4186 | -15.6546
9.077 | -34.419 | -15.6641
9.0831 | -34.4194 | -15.6737
9.0892 | -34.4198 | -15.6833
9.0953 | -34.4202 | -15.6927
9.1015 | -34.4206 | -15.7023
9.1076 | -34.421 | -15.7117
9.1137 | -34.4214 | -15.7213
9.1198 | -34.4217 | -15.7309
9.1259 | -34.4221 | -15.7405
9.132 | -34.4225 | -15.7499
9.1381 | -34.4229 | -15.7594
9.1442 | -34.4233 | -15.7687
9.1503 | -34.4237 | -15.7784
9.1564 | -34.4241 | -15.7878
9.1625 | -34.4245 | -15.7973
9.1686 | -34.4248 | -15.8067
9.1747 | -34.4252 | -15.8163
9.1808 | -34.4256 | -15.8256
9.1869 | -34.426 | -15.8351
9.193 | -34.4264 | -15.8446
9.1991 | -34.4268 | -15.8542
9.2052 | -34.4272 | -15.8638
9.2113 | -34.4276 | -15.8732
9.2174 | -34.428 | -15.8825
9.2235 | -34.4284 | -15.8919
9.2296 | -34.4288 | -15.9014
9.2357 | -34.4292 | -15.9109
9.2419 | -34.4296 | -15.9204
9.248 | -34.43 | -15.9296
9.2541 | -34.4304 | -15.939
9.2602 | -34.4308 | -15.9485
9.2663 | -34.4312 | -15.9579
9.2724 | -34.4315 | -15.9674
9.2785 | -34.4319 | -15.9769
9.2846 | -34.4323 | -15.9863
9.2907 | -34.4327 | -15.9958
9.2968 | -34.4331 | -16.0052
9.3029 | -34.4335 | -16.0148
9.309 | -34.4339 | -16.0242
9.3151 | -34.4343 | -16.0338
9.3212 | -34.4347 | -16.043
9.3273 | -34.435 | -16.0524
9.3334 | -34.4354 | -16.0619
9.3395 | -34.4358 | -16.0713
9.3456 | -34.4362 | -16.0808
9.3517 | -34.4366 | -16.0902
9.3578 | -34.437 | -16.0995
9.3639 | -34.4374 | -16.1089
9.37 | -34.4378 | -16.1184
9.3761 | -34.4382 | -16.1277
9.3822 | -34.4385 | -16.1371
9.3884 | -34.4389 | -16.1464
9.3945 | -34.4393 | -16.1558
9.4006 | -34.4397 | -16.1652
9.4067 | -34.4401 | -16.1747
9.4128 | -34.4405 | -16.1841
9.4189 | -34.4409 | -16.1935
9.425 | -34.4413 | -16.2028
9.4311 | -34.4417 | -16.2123
9.4372 | -34.4421 | -16.2216
9.4433 | -34.4425 | -16.231
9.4494 | -34.4429 | -16.2404
9.4555 | -34.4433 | -16.2497
9.4616 | -34.4437 | -16.259
9.4677 | -34.4441 | -16.2684
9.4738 | -34.4445 | -16.2778
9.4799 | -34.4449 | -16.2872
9.486 | -34.4453 | -16.2967
9.4921 | -34.4457 | -16.3061
9.4982 | -34.4461 | -16.3153
9.5043 | -34.4465 | -16.3246
9.5104 | -34.4469 | -16.334
9.5165 | -34.4473 | -16.3434
9.5226 | -34.4477 | -16.3529
9.5288 | -34.448 | -16.3621
9.5349 | -34.4484 | -16.3715
9.541 | -34.4488 | -16.381
9.5471 | -34.4492 | -16.3904
9.5532 | -34.4496 | -16.3997
9.5593 | -34.45 | -16.409
9.5654 | -34.4504 | -16.4183
9.5715 | -34.4508 | -16.4277
9.5776 | -34.4512 | -16.4371
9.5837 | -34.4516 | -16.4466
9.5898 | -34.452 | -16.4561
9.5959 | -34.4524 | -16.4653
9.602 | -34.4528 | -16.4747
9.6081 | -34.4532 | -16.484
9.6142 | -34.4536 | -16.4934
9.6203 | -34.454 | -16.5027
9.6264 | -34.4544 | -16.5122
9.6325 | -34.4548 | -16.5216
9.6386 | -34.4552 | -16.5308
9.6447 | -34.4556 | -16.5401
9.6508 | -34.456 | -16.5495
9.6569 | -34.4564 | -16.5589
9.663 | -34.4568 | -16.5681
9.6691 | -34.4572 | -16.5774
9.6753 | -34.4576 | -16.5868
9.6814 | -34.458 | -16.5963
9.6875 | -34.4584 | -16.6056
9.6936 | -34.4588 | -16.615
9.6997 | -34.4592 | -16.6244
9.7058 | -34.4596 | -16.6335
9.7119 | -34.46 | -16.6428
9.718 | -34.4604 | -16.6521
9.7241 | -34.4608 | -16.6614
9.7302 | -34.4612 | -16.6709
9.7363 | -34.4615 | -16.6802
9.7424 | -34.462 | -16.6895
9.7485 | -34.4624 | -16.6989
9.7546 | -34.4628 | -16.7083
9.7607 | -34.4632 | -16.7176
9.7668 | -34.4636 | -16.7269
9.7729 | -34.4639 | -16.7363
9.779 | -34.4644 | -16.7455
9.7851 | -34.4648 | -16.7548
9.7912 | -34.4652 | -16.7641
9.7973 | -34.4656 | -16.7735
9.8034 | -34.466 | -16.7828
9.8095 | -34.4664 | -16.7921
9.8157 | -34.4668 | -16.8013
9.8218 | -34.4672 | -16.8105
9.8279 | -34.4676 | -16.8199
9.834 | -34.468 | -16.8292
9.8401 | -34.4684 | -16.8385
9.8462 | -34.4688 | -16.8477
9.8523 | -34.4692 | -16.8571
9.8584 | -34.4696 | -16.8663
9.8645 | -34.47 | -16.8756
9.8706 | -34.4704 | -16.8849
9.8767 | -34.4708 | -16.8942
9.8828 | -34.4712 | -16.9035
9.8889 | -34.4716 | -16.9129
9.895 | -34.472 | -16.9222
9.9011 | -34.4724 | -16.9314
9.9072 | -34.4728 | -16.9407
9.9133 | -34.4732 | -16.9502
9.9194 | -34.4736 | -16.9594
9.9255 | -34.474 | -16.9686
9.9316 | -34.4745 | -16.9779
9.9377 | -34.4749 | -16.9873
9.9438 | -34.4753 | -16.9966
9.9499 | -34.4757 | -17.006
9.956 | -34.4761 | -17.0152
9.9622 | -34.4765 | -17.0245
9.9683 | -34.4769 | -17.0337
9.9744 | -34.4773 | -17.0431
9.9805 | -34.4777 | -17.0523
9.9866 | -34.4781 | -17.0615
9.9927 | -34.4785 | -17.0709
9.9988 | -34.479 | -17.0801
10.0049 | -34.4794 | -17.0894
10.011 | -34.4798 | -17.0987
10.0171 | -34.4802 | -17.108
10.0232 | -34.4806 | -17.1172
10.0293 | -34.481 | -17.1265
10.0354 | -34.4814 | -17.1356
10.0415 | -34.4818 | -17.1449
10.0476 | -34.4822 | -17.1542
10.0537 | -34.4826 | -17.1636
10.0598 | -34.4831 | -17.1729
10.0659 | -34.4835 | -17.1822
10.072 | -34.4839 | -17.1914
10.0781 | -34.4843 | -17.2007
10.0842 | -34.4847 | -17.2099
10.0903 | -34.4851 | -17.2192
10.0964 | -34.4855 | -17.2285
10.1026 | -34.4859 | -17.2378
10.1087 | -34.4864 | -17.2472
10.1148 | -34.4868 | -17.2566
10.1209 | -34.4872 | -17.2658
10.127 | -34.4876 | -17.2751
10.1331 | -34.488 | -17.2843
10.1392 | -34.4884 | -17.2937
10.1453 | -34.4888 | -17.3029
10.1514 | -34.4892 | -17.3122
10.1575 | -34.4896 | -17.3216
10.1636 | -34.49 | -17.3307
10.1697 | -34.4904 | -17.34
10.1758 | -34.4908 | -17.3493
10.1819 | -34.4912 | -17.3586
10.188 | -34.4916 | -17.3679
10.1941 | -34.4921 | -17.3772
10.2002 | -34.4925 | -17.3865
10.2063 | -34.4929 | -17.3958
10.2124 | -34.4933 | -17.405
10.2185 | -34.4937 | -17.4142
10.2246 | -34.4941 | -17.4235
10.2307 | -34.4945 | -17.4328
10.2368 | -34.4949 | -17.4419
10.2429 | -34.4953 | -17.4512
10.2491 | -34.4957 | -17.4605
10.2552 | -34.4962 | -17.4698
10.2613 | -34.4966 | -17.479
10.2674 | -34.497 | -17.4883
10.2735 | -34.4974 | -17.4975
10.2796 | -34.4978 | -17.5067
10.2857 | -34.4982 | -17.516
10.2918 | -34.4987 | -17.5252
10.2979 | -34.4991 | -17.5344
10.304 | -34.4995 | -17.5435
10.3101 | -34.4999 | -17.5529
10.3162 | -34.5004 | -17.5621
10.3223 | -34.5007 | -17.5713
10.3284 | -34.5012 | -17.5805
10.3345 | -34.5016 | -17.5899
10.3406 | -34.502 | -17.5992
10.3467 | -34.5024 | -17.6085
10.3528 | -34.5028 | -17.6178
10.3589 | -34.5032 | -17.627
10.365 | -34.5037 | -17.6362
10.3711 | -34.5041 | -17.6454
10.3772 | -34.5045 | -17.6545
10.3833 | -34.5049 | -17.6638
10.3895 | -34.5054 | -17.6729
10.3956 | -34.5058 | -17.6822
10.4017 | -34.5062 | -17.6915
10.4078 | -34.5066 | -17.7007
10.4139 | -34.5071 | -17.7099
10.42 | -34.5075 | -17.7192
10.4261 | -34.5079 | -17.7284
10.4322 | -34.5083 | -17.7377
10.4383 | -34.5087 | -17.7469
10.4444 | -34.5092 | -17.7561
10.4505 | -34.5096 | -17.7653
10.4566 | -34.51 | -17.7746
10.4627 | -34.5104 | -17.7839
10.4688 | -34.5108 | -17.7931
10.4749 | -34.5112 | -17.8023
10.481 | -34.5117 | -17.8115
10.4871 | -34.5121 | -17.8208
10.4932 | -34.5125 | -17.83
10.4993 | -34.5129 | -17.8392
10.5054 | -34.5133 | -17.8485
10.5115 | -34.5137 | -17.8577
10.5176 | -34.5141 | -17.8668
10.5237 | -34.5146 | -17.8761
10.5298 | -34.515 | -17.8853
10.536 | -34.5154 | -17.8945
10.5421 | -34.5158 | -17.9037
10.5482 | -34.5162 | -17.913
10.5543 | -34.5167 | -17.9222
10.5604 | -34.5171 | -17.9315
10.5665 | -34.5175 | -17.9406
10.5726 | -34.5179 | -17.9499
10.5787 | -34.5183 | -17.9591
10.5848 | -34.5187 | -17.9683
10.5909 | -34.5192 | -17.9775
10.597 | -34.5196 | -17.9868
10.6031 | -34.52 | -17.9961
10.6092 | -34.5204 | -18.0054
10.6153 | -34.5208 | -18.0147
10.6214 | -34.5213 | -18.0237
10.6275 | -34.5217 | -18.0331
10.6336 | -34.5221 | -18.0423
10.6397 | -34.5226 | -18.0516
10.6458 | -34.523 | -18.0608
10.6519 | -34.5234 | -18.0701
10.658 | -34.5238 | -18.0793
10.6641 | -34.5243 | -18.0884
10.6702 | -34.5247 | -18.0977
10.6764 | -34.5251 | -18.107
10.6825 | -34.5256 | -18.1162
10.6886 | -34.526 | -18.1254
10.6947 | -34.5264 | -18.1345
10.7008 | -34.5268 | -18.1437
10.7069 | -34.5273 | -18.1529
10.713 | -34.5277 | -18.1622
10.7191 | -34.5281 | -18.1713
10.7252 | -34.5285 | -18.1805
10.7313 | -34.5289 | -18.1898
10.7374 | -34.5293 | -18.199
10.7435 | -34.5298 | -18.2083
10.7496 | -34.5302 | -18.2175
10.7557 | -34.5306 | -18.2267
10.7618 | -34.531 | -18.2358
10.7679 | -34.5315 | -18.245
10.774 | -34.5319 | -18.2542
10.7801 | -34.5323 | -18.2635
10.7862 | -34.5327 | -18.2727
10.7923 | -34.5332 | -18.282
10.7984 | -34.5336 | -18.2911
10.8045 | -34.534 | -18.3003
10.8106 | -34.5344 | -18.3095
10.8168 | -34.5349 | -18.3189
10.8229 | -34.5353 | -18.3281
10.829 | -34.5357 | -18.3373
10.8351 | -34.5361 | -18.3465
10.8412 | -34.5365 | -18.3558
10.8473 | -34.537 | -18.3649
10.8534 | -34.5374 | -18.3741
10.8595 | -34.5378 | -18.3833
10.8656 | -34.5383 | -18.3925
10.8717 | -34.5387 | -18.4017
10.8778 | -34.5391 | -18.4109
10.8839 | -34.5395 | -18.4201
10.89 | -34.54 | -18.4294
10.8961 | -34.5404 | -18.4385
10.9022 | -34.5408 | -18.4478
10.9083 | -34.5412 | -18.4569
10.9144 | -34.5416 | -18.4662
10.9205 | -34.5421 | -18.4753
10.9266 | -34.5425 | -18.4845
10.9327 | -34.5429 | -18.4936
10.9388 | -34.5434 | -18.5027
10.9449 | -34.5438 | -18.5119
10.951 | -34.5442 | -18.5211
10.9571 | -34.5446 | -18.5302
10.9633 | -34.545 | -18.5395
10.9694 | -34.5455 | -18.5488
10.9755 | -34.5459 | -18.5579
10.9816 | -34.5463 | -18.567
10.9877 | -34.5468 | -18.5762
10.9938 | -34.5472 | -18.5854
10.9999 | -34.5476 | -18.5945
11.006 | -34.5481 | -18.6037
11.0121 | -34.5485 | -18.6129
11.0182 | -34.5489 | -18.622
11.0243 | -34.5493 | -18.6312
11.0304 | -34.5498 | -18.6404
11.0365 | -34.5502 | -18.6495
11.0426 | -34.5506 | -18.6586
11.0487 | -34.551 | -18.6677
11.0548 | -34.5515 | -18.6769
11.0609 | -34.5519 | -18.6861
11.067 | -34.5523 | -18.6953
11.0731 | -34.5527 | -18.7044
11.0792 | -34.5531 | -18.7135
11.0853 | -34.5536 | -18.7227
11.0914 | -34.554 | -18.7319
11.0975 | -34.5544 | -18.7411
11.1037 | -34.5548 | -18.7503
11.1098 | -34.5553 | -18.7594
11.1159 | -34.5557 | -18.7686
11.122 | -34.5561 | -18.7778
11.1281 | -34.5566 | -18.7869
11.1342 | -34.557 | -18.796
11.1403 | -34.5574 | -18.8052
11.1464 | -34.5579 | -18.8143
11.1525 | -34.5583 | -18.8235
11.1586 | -34.5587 | -18.8326
11.1647 | -34.5592 | -18.8417
11.1708 | -34.5596 | -18.8509
11.1769 | -34.56 | -18.8602
11.183 | -34.5605 | -18.8692
11.1891 | -34.5609 | -18.8784
11.1952 | -34.5613 | -18.8875
11.2013 | -34.5617 | -18.8966
11.2074 | -34.5622 | -18.9059
11.2135 | -34.5626 | -18.9151
11.2196 | -34.5631 | -18.9242
11.2257 | -34.5635 | -18.9333
11.2318 | -34.5639 | -18.9424
11.2379 | -34.5643 | -18.9516
11.244 | -34.5648 | -18.9607
11.2502 | -34.5652 | -18.9699
11.2563 | -34.5656 | -18.9789
11.2624 | -34.5661 | -18.9882
11.2685 | -34.5665 | -18.9973
11.2746 | -34.5669 | -19.0065
11.2807 | -34.5674 | -19.0156
11.2868 | -34.5678 | -19.0248
11.2929 | -34.5682 | -19.0338
11.299 | -34.5687 | -19.0428
11.3051 | -34.5691 | -19.0519
11.3112 | -34.5696 | -19.0611
11.3173 | -34.57 | -19.0701
11.3234 | -34.5704 | -19.0791
11.3295 | -34.5709 | -19.0881
11.3356 | -34.5713 | -19.0971
11.3417 | -34.5717 | -19.1063
11.3478 | -34.5722 | -19.1156
11.3539 | -34.5726 | -19.1247
11.36 | -34.573 | -19.1337
11.3661 | -34.5735 | -19.1427
11.3722 | -34.5739 | -19.1517
11.3783 | -34.5744 | -19.1608
11.3844 | -34.5748 | -19.17
11.3906 | -34.5752 | -19.179
11.3967 | -34.5757 | -19.1881
11.4028 | -34.5761 | -19.1972
11.4089 | -34.5765 | -19.2064
11.415 | -34.577 | -19.2155
11.4211 | -34.5774 | -19.2246
11.4272 | -34.5779 | -19.2337
11.4333 | -34.5783 | -19.2426
11.4394 | -34.5788 | -19.2518
11.4455 | -34.5792 | -19.261
11.4516 | -34.5796 | -19.2701
11.4577 | -34.5801 | -19.2792
11.4638 | -34.5805 | -19.2882
11.4699 | -34.581 | -19.2973
11.476 | -34.5814 | -19.3063
11.4821 | -34.5819 | -19.3154
11.4882 | -34.5823 | -19.3245
11.4943 | -34.5828 | -19.3337
11.5004 | -34.5832 | -19.3426
11.5065 | -34.5837 | -19.3516
11.5126 | -34.5841 | -19.3607
11.5187 | -34.5845 | -19.3698
11.5248 | -34.585 | -19.3788
11.5309 | -34.5854 | -19.3879
11.5371 | -34.5859 | -19.3969
11.5432 | -34.5863 | -19.406
11.5493 | -34.5868 | -19.4151
11.5554 | -34.5872 | -19.4242
11.5615 | -34.5877 | -19.4333
11.5676 | -34.5881 | -19.4424
11.5737 | -34.5886 | -19.4515
11.5798 | -34.589 | -19.4605
11.5859 | -34.5894 | -19.4695
11.592 | -34.5899 | -19.4787
11.5981 | -34.5903 | -19.4878
11.6042 | -34.5908 | -19.4968
11.6103 | -34.5912 | -19.5059
11.6164 | -34.5916 | -19.5149
11.6225 | -34.5921 | -19.524
11.6286 | -34.5925 | -19.5331
11.6347 | -34.5929 | -19.5421
11.6408 | -34.5934 | -19.5511
11.6469 | -34.5939 | -19.56
11.653 | -34.5943 | -19.5691
11.6591 | -34.5947 | -19.5782
11.6652 | -34.5951 | -19.5873
11.6713 | -34.5956 | -19.5964
11.6775 | -34.596 | -19.6054
11.6836 | -34.5965 | -19.6144
11.6897 | -34.5969 | -19.6234
11.6958 | -34.5974 | -19.6325
11.7019 | -34.5978 | -19.6414
11.708 | -34.5983 | -19.6505
11.7141 | -34.5988 | -19.6593
11.7202 | -34.5992 | -19.6684
11.7263 | -34.5996 | -19.6775
11.7324 | -34.6001 | -19.6865
11.7385 | -34.6005 | -19.6956
11.7446 | -34.601 | -19.7048
11.7507 | -34.6014 | -19.7138
11.7568 | -34.6018 | -19.7228
11.7629 | -34.6023 | -19.7319
11.769 | -34.6027 | -19.7409
11.7751 | -34.6032 | -19.7499
11.7812 | -34.6036 | -19.7588
11.7873 | -34.6041 | -19.7678
11.7934 | -34.6045 | -19.7769
11.7995 | -34.605 | -19.786
11.8056 | -34.6054 | -19.795
11.8117 | -34.6059 | -19.8041
11.8178 | -34.6063 | -19.8129
11.824 | -34.6068 | -19.8219
11.8301 | -34.6072 | -19.8309
11.8362 | -34.6076 | -19.84
11.8423 | -34.6081 | -19.8491
11.8484 | -34.6085 | -19.8581
11.8545 | -34.609 | -19.867
11.8606 | -34.6094 | -19.876
11.8667 | -34.6099 | -19.8849
11.8728 | -34.6103 | -19.8938
11.8789 | -34.6107 | -19.9029
11.885 | -34.6112 | -19.9119
11.8911 | -34.6117 | -19.9207
11.8972 | -34.6121 | -19.9297
11.9033 | -34.6125 | -19.9387
11.9094 | -34.613 | -19.9477
11.9155 | -34.6134 | -19.9567
11.9216 | -34.6139 | -19.9655
11.9277 | -34.6143 | -19.9745
11.9338 | -34.6148 | -19.9835
11.9399 | -34.6152 | -19.9925
11.946 | -34.6157 | -20.0013
11.9521 | -34.6161 | -20.0103
11.9582 | -34.6165 | -20.0192
11.9644 | -34.617 | -20.0282
11.9705 | -34.6174 | -20.0371
11.9766 | -34.6179 | -20.046
11.9827 | -34.6183 | -20.0549
11.9888 | -34.6188 | -20.0637
11.9949 | -34.6192 | -20.0726
12.001 | -34.6197 | -20.0816
12.0071 | -34.6201 | -20.0905
12.0132 | -34.6205 | -20.0995
12.0193 | -34.621 | -20.1084
12.0254 | -34.6214 | -20.1174
12.0315 | -34.6219 | -20.1261
12.0376 | -34.6223 | -20.1351
12.0437 | -34.6228 | -20.1442
12.0498 | -34.6232 | -20.153
12.0559 | -34.6237 | -20.162
12.062 | -34.6241 | -20.1708
12.0681 | -34.6246 | -20.1797
12.0742 | -34.625 | -20.1886
12.0803 | -34.6255 | -20.1977
12.0864 | -34.6259 | -20.2066
12.0925 | -34.6264 | -20.2155
12.0986 | -34.6269 | -20.2243
12.1047 | -34.6273 | -20.2333
12.1109 | -34.6277 | -20.2422
12.117 | -34.6282 | -20.2511
12.1231 | -34.6286 | -20.26
12.1292 | -34.6291 | -20.269
12.1353 | -34.6296 | -20.2777
12.1414 | -34.63 | -20.2865
12.1475 | -34.6305 | -20.2955
12.1536 | -34.6309 | -20.3044
12.1597 | -34.6313 | -20.3132
12.1658 | -34.6318 | -20.3221
12.1719 | -34.6323 | -20.331
12.178 | -34.6327 | -20.3399
12.1841 | -34.6332 | -20.3488
12.1902 | -34.6336 | -20.3576
12.1963 | -34.6341 | -20.3665
12.2024 | -34.6345 | -20.3751
12.2085 | -34.635 | -20.384
12.2146 | -34.6354 | -20.3929
12.2207 | -34.6359 | -20.4018
12.2268 | -34.6363 | -20.4107
12.2329 | -34.6368 | -20.4196
12.239 | -34.6372 | -20.4283
12.2451 | -34.6377 | -20.4371
12.2513 | -34.6381 | -20.446
12.2574 | -34.6386 | -20.4549
12.2635 | -34.639 | -20.4638
12.2696 | -34.6395 | -20.4727
12.2757 | -34.6399 | -20.4814
12.2818 | -34.6404 | -20.4903
12.2879 | -34.6408 | -20.4991
12.294 | -34.6413 | -20.508
12.3001 | -34.6417 | -20.5169
12.3062 | -34.6422 | -20.5257
12.3123 | -34.6426 | -20.5344
12.3184 | -34.6431 | -20.5433
12.3245 | -34.6435 | -20.5522
12.3306 | -34.6439 | -20.5611
12.3367 | -34.6444 | -20.5698
12.3428 | -34.6448 | -20.5787
12.3489 | -34.6453 | -20.5873
12.355 | -34.6457 | -20.5962
12.3611 | -34.6462 | -20.605
12.3672 | -34.6466 | -20.6139
12.3733 | -34.6471 | -20.6227
12.3794 | -34.6475 | -20.6313
12.3855 | -34.648 | -20.6401
12.3916 | -34.6484 | -20.649
12.3978 | -34.6489 | -20.6578
12.4039 | -34.6493 | -20.6667
12.41 | -34.6498 | -20.6755
12.4161 | -34.6502 | -20.6841
12.4222 | -34.6507 | -20.6928
12.4283 | -34.6511 | -20.7017
12.4344 | -34.6516 | -20.7104
12.4405 | -34.652 | -20.7192
12.4466 | -34.6524 | -20.728
12.4527 | -34.6529 | -20.7368
12.4588 | -34.6534 | -20.7456
12.4649 | -34.6538 | -20.7543
12.471 | -34.6543 | -20.7631
12.4771 | -34.6547 | -20.772
12.4832 | -34.6552 | -20.7809
12.4893 | -34.6556 | -20.7896
12.4954 | -34.6561 | -20.7983
12.5015 | -34.6565 | -20.8072
12.5076 | -34.657 | -20.816
12.5137 | -34.6574 | -20.8247
12.5198 | -34.6579 | -20.8334
12.5259 | -34.6584 | -20.8422
12.532 | -34.6588 | -20.851
12.5382 | -34.6593 | -20.8597
12.5443 | -34.6597 | -20.8686
12.5504 | -34.6602 | -20.8774
12.5565 | -34.6606 | -20.8859
12.5626 | -34.6611 | -20.8947
12.5687 | -34.6616 | -20.9034
12.5748 | -34.662 | -20.9122
12.5809 | -34.6625 | -20.9209
12.587 | -34.6629 | -20.9296
12.5931 | -34.6634 | -20.9382
12.5992 | -34.6639 | -20.9468
12.6053 | -34.6643 | -20.9557
12.6114 | -34.6648 | -20.9645
12.6175 | -34.6652 | -20.9732
12.6236 | -34.6657 | -20.982
12.6297 | -34.6661 | -20.9905
12.6358 | -34.6666 | -20.9993
12.6419 | -34.667 | -21.008
12.648 | -34.6675 | -21.0168
12.6541 | -34.6679 | -21.0256
12.6602 | -34.6684 | -21.034
12.6663 | -34.6688 | -21.0427
12.6724 | -34.6693 | -21.0513
12.6785 | -34.6698 | -21.06
12.6847 | -34.6702 | -21.0688
12.6908 | -34.6707 | -21.0774
12.6969 | -34.6711 | -21.0859
12.703 | -34.6716 | -21.0947
12.7091 | -34.672 | -21.1034
12.7152 | -34.6725 | -21.1122
12.7213 | -34.6729 | -21.121
12.7274 | -34.6733 | -21.1296
12.7335 | -34.6738 | -21.1382
12.7396 | -34.6742 | -21.1469
12.7457 | -34.6747 | -21.1557
12.7518 | -34.6752 | -21.1644
12.7579 | -34.6756 | -21.1731
12.764 | -34.6761 | -21.1818
12.7701 | -34.6766 | -21.1904
12.7762 | -34.677 | -21.1991
12.7823 | -34.6775 | -21.2076
12.7884 | -34.6779 | -21.2162
12.7945 | -34.6784 | -21.225
12.8006 | -34.6788 | -21.2335
12.8067 | -34.6793 | -21.2421
12.8128 | -34.6797 | -21.2509
12.8189 | -34.6802 | -21.2596
12.8251 | -34.6806 | -21.2682
12.8312 | -34.6811 | -21.2769
12.8373 | -34.6816 | -21.2853
12.8434 | -34.682 | -21.2939
12.8495 | -34.6824 | -21.3025
12.8556 | -34.6829 | -21.3112
12.8617 | -34.6833 | -21.3196
12.8678 | -34.6838 | -21.3283
12.8739 | -34.6843 | -21.3368
12.88 | -34.6847 | -21.3455
12.8861 | -34.6852 | -21.3541
12.8922 | -34.6856 | -21.3628
12.8983 | -34.6861 | -21.3715
12.9044 | -34.6865 | -21.38
12.9105 | -34.687 | -21.3884
12.9166 | -34.6874 | -21.397
12.9227 | -34.6879 | -21.4057
12.9288 | -34.6883 | -21.4144
12.9349 | -34.6888 | -21.4231
12.941 | -34.6892 | -21.4316
12.9471 | -34.6897 | -21.4402
12.9532 | -34.6902 | -21.4489
12.9593 | -34.6906 | -21.4577
12.9654 | -34.691 | -21.4664
12.9716 | -34.6915 | -21.475
12.9777 | -34.692 | -21.4833
12.9838 | -34.6925 | -21.4921
12.9899 | -34.6929 | -21.5007
12.996 | -34.6934 | -21.5092
13.0021 | -34.6938 | -21.5177
13.0082 | -34.6943 | -21.5261
13.0143 | -34.6948 | -21.5343
13.0204 | -34.6952 | -21.543
13.0265 | -34.6956 | -21.5514
13.0326 | -34.6961 | -21.5603
13.0387 | -34.6966 | -21.5688
13.0448 | -34.697 | -21.5774
13.0509 | -34.6975 | -21.5857
13.057 | -34.6979 | -21.5943
13.0631 | -34.6984 | -21.6029
13.0692 | -34.6988 | -21.6115
13.0753 | -34.6992 | -21.6199
13.0814 | -34.6997 | -21.6286
13.0875 | -34.7002 | -21.6367
13.0936 | -34.7006 | -21.6453
13.0997 | -34.7011 | -21.6539
13.1058 | -34.7016 | -21.6623
13.112 | -34.702 | -21.6708
13.1181 | -34.7025 | -21.679
13.1242 | -34.703 | -21.6876
13.1303 | -34.7034 | -21.6961
13.1364 | -34.7039 | -21.7049
13.1425 | -34.7044 | -21.7134
13.1486 | -34.7048 | -21.722
13.1547 | -34.7053 | -21.7305
13.1608 | -34.7057 | -21.7394
13.1669 | -34.7062 | -21.7479
13.173 | -34.7067 | -21.7564
13.1791 | -34.7072 | -21.7651
13.1852 | -34.7076 | -21.7736
13.1913 | -34.7081 | -21.7818
13.1974 | -34.7085 | -21.7905
13.2035 | -34.709 | -21.7988
13.2096 | -34.7094 | -21.8072
13.2157 | -34.7098 | -21.8158
13.2218 | -34.7102 | -21.8243
13.2279 | -34.7107 | -21.8324
13.234 | -34.7111 | -21.841
13.2401 | -34.7116 | -21.8495
13.2462 | -34.712 | -21.8581
13.2524 | -34.7125 | -21.8667
13.2585 | -34.7129 | -21.8751
13.2646 | -34.7134 | -21.8839
13.2707 | -34.7138 | -21.8924
13.2768 | -34.7143 | -21.9007
13.2829 | -34.7147 | -21.909
13.289 | -34.7151 | -21.9174
13.2951 | -34.7156 | -21.9258
13.3012 | -34.716 | -21.9343
13.3073 | -34.7165 | -21.9428
13.3134 | -34.717 | -21.9512
13.3195 | -34.7174 | -21.9599
13.3256 | -34.7178 | -21.9683
13.3317 | -34.7183 | -21.9767
13.3378 | -34.7187 | -21.9852
13.3439 | -34.7192 | -21.9939
13.35 | -34.7197 | -22.0025
13.3561 | -34.7201 | -22.0112
13.3622 | -34.7206 | -22.0198
13.3683 | -34.721 | -22.0281
13.3744 | -34.7215 | -22.0367
13.3805 | -34.7219 | -22.0452
13.3866 | -34.7224 | -22.0538
13.3927 | -34.7229 | -22.0623
13.3989 | -34.7233 | -22.0706
13.405 | -34.7238 | -22.0791
13.4111 | -34.7242 | -22.0875
13.4172 | -34.7247 | -22.0959
13.4233 | -34.7251 | -22.1045
13.4294 | -34.7256 | -22.1131
13.4355 | -34.726 | -22.1213
13.4416 | -34.7265 | -22.1298
13.4477 | -34.7269 | -22.1382
13.4538 | -34.7274 | -22.1466
13.4599 | -34.7278 | -22.1552
13.466 | -34.7283 | -22.1636
13.4721 | -34.7288 | -22.172
13.4782 | -34.7292 | -22.1804
13.4843 | -34.7296 | -22.1886
13.4904 | -34.7301 | -22.1971
13.4965 | -34.7305 | -22.2057
13.5026 | -34.731 | -22.2141
13.5087 | -34.7315 | -22.2225
13.5148 | -34.732 | -22.2308
13.5209 | -34.7325 | -22.2393
13.527 | -34.7329 | -22.2479
13.5331 | -34.7334 | -22.2562
13.5393 | -34.7338 | -22.2642
13.5454 | -34.7343 | -22.2723
13.5515 | -34.7348 | -22.2808
13.5576 | -34.7352 | -22.2893
13.5637 | -34.7357 | -22.2977
13.5698 | -34.7362 | -22.306
13.5759 | -34.7366 | -22.3141
13.582 | -34.7371 | -22.3225
13.5881 | -34.7376 | -22.3309
13.5942 | -34.738 | -22.3393
13.6003 | -34.7385 | -22.3479
13.6064 | -34.7389 | -22.3566
13.6125 | -34.7394 | -22.3649
13.6186 | -34.7399 | -22.3733
13.6247 | -34.7403 | -22.3817
13.6308 | -34.7408 | -22.3901
13.6369 | -34.7413 | -22.3987
13.643 | -34.7417 | -22.4071
13.6491 | -34.7422 | -22.4153
13.6552 | -34.7426 | -22.4237
13.6613 | -34.7431 | -22.4322
13.6674 | -34.7436 | -22.4406
13.6735 | -34.744 | -22.4491
13.6796 | -34.7445 | -22.4575
13.6858 | -34.7449 | -22.4657
13.6919 | -34.7454 | -22.4739
13.698 | -34.7459 | -22.4822
13.7041 | -34.7464 | -22.4905
13.7102 | -34.7468 | -22.4992
13.7163 | -34.7473 | -22.5073
13.7224 | -34.7477 | -22.5157
13.7285 | -34.7482 | -22.524
13.7346 | -34.7486 | -22.5324
13.7407 | -34.7491 | -22.5407
13.7468 | -34.7495 | -22.5491
13.7529 | -34.75 | -22.5574
13.759 | -34.7505 | -22.5659
13.7651 | -34.751 | -22.5741
13.7712 | -34.7514 | -22.5825
13.7773 | -34.7519 | -22.5909
13.7834 | -34.7524 | -22.5992
13.7895 | -34.7528 | -22.6072
13.7956 | -34.7533 | -22.6156
13.8017 | -34.7538 | -22.6241
13.8078 | -34.7543 | -22.6326
13.8139 | -34.7547 | -22.641
13.82 | -34.7552 | -22.6494
13.8262 | -34.7557 | -22.6578
13.8323 | -34.7562 | -22.666
13.8384 | -34.7566 | -22.6742
13.8445 | -34.7571 | -22.6826
13.8506 | -34.7575 | -22.6909
13.8567 | -34.758 | -22.699
13.8628 | -34.7585 | -22.7074
13.8689 | -34.7589 | -22.7157
13.875 | -34.7594 | -22.7241
13.8811 | -34.7598 | -22.7326
13.8872 | -34.7603 | -22.7409
13.8933 | -34.7608 | -22.7491
13.8994 | -34.7612 | -22.7573
13.9055 | -34.7617 | -22.7655
13.9116 | -34.7622 | -22.7738
13.9177 | -34.7626 | -22.7821
13.9238 | -34.7631 | -22.7903
13.9299 | -34.7635 | -22.7985
13.936 | -34.764 | -22.8069
13.9421 | -34.7645 | -22.8152
13.9482 | -34.7649 | -22.8235
13.9543 | -34.7654 | -22.8319
13.9604 | -34.7658 | -22.8403
13.9665 | -34.7663 | -22.8485
13.9727 | -34.7668 | -22.857
13.9788 | -34.7672 | -22.865
13.9849 | -34.7677 | -22.8733
13.991 | -34.7682 | -22.8815
13.9971 | -34.7686 | -22.8899
14.0032 | -34.769 | -22.8983
14.0093 | -34.7695 | -22.9067
14.0154 | -34.7699 | -22.9149
14.0215 | -34.7704 | -22.9231
14.0276 | -34.7708 | -22.9314
14.0337 | -34.7714 | -22.9395
14.0398 | -34.7718 | -22.9479
14.0459 | -34.7723 | -22.9561
14.052 | -34.7728 | -22.9643
14.0581 | -34.7732 | -22.9728
14.0642 | -34.7737 | -22.9811
14.0703 | -34.7742 | -22.9894
14.0764 | -34.7746 | -22.9976
14.0825 | -34.7751 | -23.006
14.0886 | -34.7755 | -23.0143
14.0947 | -34.776 | -23.0224
14.1008 | -34.7765 | -23.0307
14.1069 | -34.777 | -23.0389
14.1131 | -34.7774 | -23.0471
14.1192 | -34.7778 | -23.0553
14.1253 | -34.7783 | -23.0634
14.1314 | -34.7788 | -23.0716
14.1375 | -34.7792 | -23.0796
14.1436 | -34.7797 | -23.0877
14.1497 | -34.7802 | -23.0961
14.1558 | -34.7806 | -23.1043
14.1619 | -34.7811 | -23.1126
14.168 | -34.7815 | -23.1208
14.1741 | -34.782 | -23.1289
14.1802 | -34.7825 | -23.1372
14.1863 | -34.783 | -23.1454
14.1924 | -34.7834 | -23.1537
14.1985 | -34.7839 | -23.1619
14.2046 | -34.7844 | -23.1703
14.2107 | -34.7848 | -23.1783
14.2168 | -34.7852 | -23.1865
14.2229 | -34.7857 | -23.1949
14.229 | -34.7862 | -23.203
14.2351 | -34.7866 | -23.2112
14.2412 | -34.7871 | -23.2194
14.2473 | -34.7875 | -23.2276
14.2534 | -34.788 | -23.236
14.2596 | -34.7884 | -23.2443
14.2657 | -34.7889 | -23.2527
14.2718 | -34.7894 | -23.2608
14.2779 | -34.7899 | -23.2689
14.284 | -34.7903 | -23.277
14.2901 | -34.7908 | -23.2853
14.2962 | -34.7913 | -23.2935
14.3023 | -34.7917 | -23.3017
14.3084 | -34.7922 | -23.3101
14.3145 | -34.7927 | -23.3181
14.3206 | -34.7932 | -23.3265
14.3267 | -34.7936 | -23.3346
14.3328 | -34.7941 | -23.3429
14.3389 | -34.7945 | -23.3511
14.345 | -34.795 | -23.3592
14.3511 | -34.7955 | -23.3672
14.3572 | -34.796 | -23.3754
14.3633 | -34.7964 | -23.3837
14.3694 | -34.7969 | -23.3917
14.3755 | -34.7974 | -23.3999
14.3816 | -34.7979 | -23.4082
14.3877 | -34.7984 | -23.4161
14.3938 | -34.7988 | -23.4244
14.4 | -34.7993 | -23.4327
14.4061 | -34.7997 | -23.4409
14.4122 | -34.8002 | -23.4492
14.4183 | -34.8007 | -23.4574
14.4244 | -34.8012 | -23.4654
14.4305 | -34.8016 | -23.4736
14.4366 | -34.8021 | -23.4817
14.4427 | -34.8026 | -23.4898
14.4488 | -34.803 | -23.498
14.4549 | -34.8035 | -23.5061
14.461 | -34.804 | -23.5143
14.4671 | -34.8045 | -23.5225
14.4732 | -34.805 | -23.5307
14.4793 | -34.8054 | -23.5388
14.4854 | -34.8059 | -23.547
14.4915 | -34.8064 | -23.555
14.4976 | -34.8068 | -23.5631
14.5037 | -34.8073 | -23.5713
14.5098 | -34.8078 | -23.5795
14.5159 | -34.8082 | -23.5877
14.522 | -34.8087 | -23.5959
14.5281 | -34.8092 | -23.6039
14.5342 | -34.8097 | -23.612
14.5403 | -34.8101 | -23.6202
14.5465 | -34.8106 | -23.6283
14.5526 | -34.8111 | -23.6366
14.5587 | -34.8116 | -23.6448
14.5648 | -34.812 | -23.6527
14.5709 | -34.8125 | -23.6608
14.577 | -34.8129 | -23.669
14.5831 | -34.8134 | -23.6772
14.5892 | -34.8139 | -23.6852
14.5953 | -34.8144 | -23.6933
14.6014 | -34.8148 | -23.7015
14.6075 | -34.8153 | -23.7097
14.6136 | -34.8158 | -23.718
14.6197 | -34.8162 | -23.7261
14.6258 | -34.8167 | -23.7343
14.6319 | -34.8172 | -23.7424
14.638 | -34.8176 | -23.7505
14.6441 | -34.8181 | -23.7588
14.6502 | -34.8186 | -23.767
14.6563 | -34.8191 | -23.7751
14.6624 | -34.8195 | -23.7833
14.6685 | -34.82 | -23.791
14.6746 | -34.8205 | -23.7992
14.6807 | -34.821 | -23.8073
14.6869 | -34.8214 | -23.8153
14.693 | -34.8219 | -23.8233
14.6991 | -34.8224 | -23.8316
14.7052 | -34.8228 | -23.8394
14.7113 | -34.8233 | -23.8474
14.7174 | -34.8238 | -23.8556
14.7235 | -34.8242 | -23.8637
14.7296 | -34.8247 | -23.8719
14.7357 | -34.8252 | -23.8798
14.7418 | -34.8257 | -23.8878
14.7479 | -34.8262 | -23.896
14.754 | -34.8266 | -23.9043
14.7601 | -34.8271 | -23.9123
14.7662 | -34.8275 | -23.9204
14.7723 | -34.828 | -23.9284
14.7784 | -34.8285 | -23.9365
14.7845 | -34.8289 | -23.9447
14.7906 | -34.8294 | -23.9527
14.7967 | -34.8299 | -23.9609
14.8028 | -34.8303 | -23.9691
14.8089 | -34.8308 | -23.9769
14.815 | -34.8313 | -23.985
14.8211 | -34.8318 | -23.993
14.8272 | -34.8322 | -24.0011
14.8334 | -34.8327 | -24.0091
14.8395 | -34.8332 | -24.0172
14.8456 | -34.8337 | -24.0251
14.8517 | -34.8341 | -24.0333
14.8578 | -34.8346 | -24.0414
14.8639 | -34.8351 | -24.0495
14.87 | -34.8355 | -24.0575
14.8761 | -34.836 | -24.0655
14.8822 | -34.8365 | -24.0737
14.8883 | -34.837 | -24.0818
14.8944 | -34.8375 | -24.0899
14.9005 | -34.8379 | -24.0979
14.9066 | -34.8384 | -24.106
14.9127 | -34.8389 | -24.114
14.9188 | -34.8394 | -24.1222
14.9249 | -34.8398 | -24.1303
14.931 | -34.8403 | -24.1384
14.9371 | -34.8408 | -24.1465
14.9432 | -34.8413 | -24.1545
14.9493 | -34.8418 | -24.1626
14.9554 | -34.8423 | -24.1707
14.9615 | -34.8427 | -24.1787
14.9676 | -34.8432 | -24.1869
14.9738 | -34.8437 | -24.1949
14.9799 | -34.8442 | -24.2029
14.986 | -34.8446 | -24.2108
14.9921 | -34.8451 | -24.2189
14.9982 | -34.8456 | -24.2269
15.0043 | -34.846 | -24.235
15.0104 | -34.8465 | -24.243
15.0165 | -34.847 | -24.251
15.0226 | -34.8475 | -24.2588
15.0287 | -34.848 | -24.267
15.0348 | -34.8485 | -24.275
15.0409 | -34.849 | -24.2829
15.047 | -34.8494 | -24.2909
15.0531 | -34.8499 | -24.2987
15.0592 | -34.8504 | -24.3068
15.0653 | -34.8509 | -24.3148
15.0714 | -34.8513 | -24.3228
15.0775 | -34.8518 | -24.3309
15.0836 | -34.8522 | -24.3388
15.0897 | -34.8527 | -24.3466
15.0958 | -34.8532 | -24.3547
15.1019 | -34.8537 | -24.3627
15.108 | -34.8541 | -24.3706
15.1141 | -34.8546 | -24.3787
15.1203 | -34.8551 | -24.3867
15.1264 | -34.8556 | -24.3945
15.1325 | -34.8561 | -24.4026
15.1386 | -34.8565 | -24.4106
15.1447 | -34.857 | -24.4185
15.1508 | -34.8575 | -24.4265
15.1569 | -34.858 | -24.4346
15.163 | -34.8584 | -24.4424
15.1691 | -34.8589 | -24.4504
15.1752 | -34.8594 | -24.4584
15.1813 | -34.8599 | -24.4665
15.1874 | -34.8603 | -24.4745
15.1935 | -34.8608 | -24.4824
15.1996 | -34.8613 | -24.4905
15.2057 | -34.8617 | -24.4984
15.2118 | -34.8622 | -24.5066
15.2179 | -34.8627 | -24.5145
15.224 | -34.8632 | -24.5226
15.2301 | -34.8637 | -24.5306
15.2362 | -34.8641 | -24.5386
15.2423 | -34.8646 | -24.5466
15.2484 | -34.8651 | -24.5545
15.2545 | -34.8656 | -24.5625
15.2607 | -34.866 | -24.5706
15.2668 | -34.8665 | -24.5785
15.2729 | -34.867 | -24.5864
15.279 | -34.8675 | -24.5945
15.2851 | -34.8679 | -24.6025
15.2912 | -34.8684 | -24.6105
15.2973 | -34.8689 | -24.6184
15.3034 | -34.8694 | -24.6262
15.3095 | -34.8699 | -24.6341
15.3156 | -34.8704 | -24.642
15.3217 | -34.8708 | -24.65
15.3278 | -34.8713 | -24.6579
15.3339 | -34.8718 | -24.6656
15.34 | -34.8722 | -24.6734
15.3461 | -34.8727 | -24.6815
15.3522 | -34.8732 | -24.6895
15.3583 | -34.8737 | -24.6976
15.3644 | -34.8741 | -24.7055
15.3705 | -34.8746 | -24.7132
15.3766 | -34.8751 | -24.7212
15.3827 | -34.8756 | -24.7291
15.3888 | -34.8761 | -24.737
15.3949 | -34.8765 | -24.745
15.401 | -34.877 | -24.7529
15.4072 | -34.8775 | -24.7608
15.4133 | -34.878 | -24.7687
15.4194 | -34.8785 | -24.7766
15.4255 | -34.879 | -24.7844
15.4316 | -34.8794 | -24.7923
15.4377 | -34.8799 | -24.8002
15.4438 | -34.8804 | -24.808
15.4499 | -34.8809 | -24.8159
15.456 | -34.8814 | -24.8239
15.4621 | -34.8818 | -24.8317
15.4682 | -34.8823 | -24.8398
15.4743 | -34.8828 | -24.8474
15.4804 | -34.8833 | -24.8555
15.4865 | -34.8837 | -24.8635
15.4926 | -34.8842 | -24.8714
15.4987 | -34.8847 | -24.8792
15.5048 | -34.8852 | -24.8872
15.5109 | -34.8857 | -24.8949
15.517 | -34.8862 | -24.9028
15.5231 | -34.8866 | -24.9108
15.5292 | -34.8871 | -24.9189
15.5353 | -34.8876 | -24.9268
15.5414 | -34.8881 | -24.9347
15.5476 | -34.8886 | -24.9425
15.5537 | -34.8891 | -24.9503
15.5598 | -34.8895 | -24.9583
15.5659 | -34.89 | -24.9662
15.572 | -34.8905 | -24.9741
15.5781 | -34.891 | -24.9819
15.5842 | -34.8915 | -24.9897
15.5903 | -34.8919 | -24.9976
15.5964 | -34.8924 | -25.0055
15.6025 | -34.8929 | -25.0134
15.6086 | -34.8934 | -25.0213
15.6147 | -34.8939 | -25.0292
15.6208 | -34.8944 | -25.0368
15.6269 | -34.8949 | -25.0447
15.633 | -34.8953 | -25.0526
15.6391 | -34.8958 | -25.0605
15.6452 | -34.8963 | -25.0683
15.6513 | -34.8968 | -25.076
15.6574 | -34.8973 | -25.0839
15.6635 | -34.8978 | -25.0918
15.6696 | -34.8982 | -25.0997
15.6757 | -34.8987 | -25.1076
15.6818 | -34.8992 | -25.1156
15.688 | -34.8997 | -25.1234
15.6941 | -34.9002 | -25.1313
15.7002 | -34.9006 | -25.1393
15.7063 | -34.9011 | -25.1472
15.7124 | -34.9016 | -25.1551
15.7185 | -34.9021 | -25.163
15.7246 | -34.9026 | -25.1707
15.7307 | -34.903 | -25.1784
15.7368 | -34.9035 | -25.1864
15.7429 | -34.904 | -25.1942
15.749 | -34.9045 | -25.2019
15.7551 | -34.9049 | -25.2098
15.7612 | -34.9055 | -25.2174
15.7673 | -34.9059 | -25.2253
15.7734 | -34.9064 | -25.2331
15.7795 | -34.9069 | -25.241
15.7856 | -34.9074 | -25.2488
15.7917 | -34.9079 | -25.2565
15.7978 | -34.9084 | -25.2645
15.8039 | -34.9089 | -25.2723
15.81 | -34.9093 | -25.2802
15.8161 | -34.9098 | -25.2881
15.8222 | -34.9103 | -25.296
15.8283 | -34.9108 | -25.3037
15.8345 | -34.9113 | -25.3114
15.8406 | -34.9118 | -25.3194
15.8467 | -34.9123 | -25.3271
15.8528 | -34.9127 | -25.3349
15.8589 | -34.9132 | -25.3428
15.865 | -34.9137 | -25.3505
15.8711 | -34.9142 | -25.3584
15.8772 | -34.9146 | -25.3661
15.8833 | -34.9151 | -25.3739
15.8894 | -34.9156 | -25.3818
15.8955 | -34.916 | -25.3897
15.9016 | -34.9165 | -25.3974
15.9077 | -34.917 | -25.4051
15.9138 | -34.9175 | -25.4131
15.9199 | -34.918 | -25.4209
15.926 | -34.9185 | -25.4288
15.9321 | -34.9189 | -25.4364
15.9382 | -34.9194 | -25.4443
15.9443 | -34.9199 | -25.4521
15.9504 | -34.9204 | -25.4599
15.9565 | -34.9209 | -25.4677
15.9626 | -34.9214 | -25.4755
15.9687 | -34.9219 | -25.483
15.9749 | -34.9223 | -25.4908
15.981 | -34.9228 | -25.4987
15.9871 | -34.9233 | -25.5065
15.9932 | -34.9238 | -25.5143
15.9993 | -34.9243 | -25.5221
16.0054 | -34.9247 | -25.5298
16.0115 | -34.9252 | -25.5375
16.0176 | -34.9257 | -25.5454
16.0237 | -34.9262 | -25.5532
16.0298 | -34.9266 | -25.5611
16.0359 | -34.9271 | -25.5689
16.042 | -34.9276 | -25.5766
16.0481 | -34.9281 | -25.5845
16.0542 | -34.9286 | -25.5923
16.0603 | -34.929 | -25.6002
16.0664 | -34.9295 | -25.6081
16.0725 | -34.93 | -25.6157
16.0786 | -34.9305 | -25.6234
16.0847 | -34.931 | -25.6312
16.0908 | -34.9315 | -25.6389
16.0969 | -34.932 | -25.6468
16.103 | -34.9324 | -25.6546
16.1091 | -34.9329 | -25.6622
16.1152 | -34.9334 | -25.6701
16.1214 | -34.9339 | -25.6779
16.1275 | -34.9344 | -25.6857
16.1336 | -34.9348 | -25.6935
16.1397 | -34.9353 | -25.7012
16.1458 | -34.9358 | -25.7088
16.1519 | -34.9363 | -25.7165
16.158 | -34.9368 | -25.7243
16.1641 | -34.9373 | -25.7321
16.1702 | -34.9378 | -25.7399
16.1763 | -34.9383 | -25.7477
16.1824 | -34.9388 | -25.7554
16.1885 | -34.9392 | -25.7632
16.1946 | -34.9397 | -25.7709
16.2007 | -34.9402 | -25.7786
16.2068 | -34.9406 | -25.7862
16.2129 | -34.9411 | -25.794
16.219 | -34.9416 | -25.8015
16.2251 | -34.942 | -25.8093
16.2312 | -34.9425 | -25.8171
16.2373 | -34.943 | -25.8248
16.2434 | -34.9435 | -25.8326
16.2495 | -34.9439 | -25.8402
16.2556 | -34.9444 | -25.848
16.2618 | -34.9449 | -25.8558
16.2679 | -34.9454 | -25.8636
16.274 | -34.9459 | -25.8714
16.2801 | -34.9463 | -25.8793
16.2862 | -34.9469 | -25.8869
16.2923 | -34.9474 | -25.8946
16.2984 | -34.9478 | -25.9024
16.3045 | -34.9483 | -25.9103
16.3106 | -34.9488 | -25.9181
16.3167 | -34.9493 | -25.9257
16.3228 | -34.9498 | -25.9332
16.3289 | -34.9502 | -25.9409
16.335 | -34.9507 | -25.9485
16.3411 | -34.9512 | -25.9562
16.3472 | -34.9517 | -25.964
16.3533 | -34.9522 | -25.9718
16.3594 | -34.9526 | -25.9795
16.3655 | -34.9531 | -25.9872
16.3716 | -34.9535 | -25.9949
16.3777 | -34.954 | -26.0026
16.3838 | -34.9545 | -26.0103
16.3899 | -34.955 | -26.0178
16.396 | -34.9555 | -26.0255
16.4021 | -34.956 | -26.0332
16.4083 | -34.9565 | -26.0409
16.4144 | -34.9569 | -26.0487
16.4205 | -34.9574 | -26.0562
16.4266 | -34.958 | -26.064
16.4327 | -34.9584 | -26.0718
16.4388 | -34.9589 | -26.0794
16.4449 | -34.9594 | -26.0871
16.451 | -34.9599 | -26.0948
16.4571 | -34.9604 | -26.1026
16.4632 | -34.9608 | -26.1099
16.4693 | -34.9613 | -26.1176
16.4754 | -34.9618 | -26.1253
16.4815 | -34.9622 | -26.1332
16.4876 | -34.9627 | -26.1408
16.4937 | -34.9632 | -26.1484
16.4998 | -34.9637 | -26.1558
16.5059 | -34.9642 | -26.1636
16.512 | -34.9647 | -26.1713
16.5181 | -34.9652 | -26.179
16.5242 | -34.9656 | -26.1867
16.5303 | -34.9662 | -26.1943
16.5364 | -34.9666 | -26.202
16.5425 | -34.9671 | -26.2097
16.5487 | -34.9676 | -26.2176
16.5548 | -34.9681 | -26.2251
16.5609 | -34.9685 | -26.2326
16.567 | -34.969 | -26.2402
16.5731 | -34.9695 | -26.2479
16.5792 | -34.97 | -26.2557
16.5853 | -34.9705 | -26.2633
16.5914 | -34.9709 | -26.271
16.5975 | -34.9714 | -26.2786
16.6036 | -34.9719 | -26.2862
16.6097 | -34.9724 | -26.294
16.6158 | -34.9729 | -26.3019
16.6219 | -34.9733 | -26.3094
16.628 | -34.9738 | -26.3171
16.6341 | -34.9743 | -26.3247
16.6402 | -34.9748 | -26.3322
16.6463 | -34.9753 | -26.34
16.6524 | -34.9757 | -26.3476
16.6585 | -34.9762 | -26.3553
16.6646 | -34.9767 | -26.3629
16.6707 | -34.9772 | -26.3707
16.6768 | -34.9777 | -26.3784
16.6829 | -34.9782 | -26.386
16.689 | -34.9787 | -26.3935
16.6952 | -34.9792 | -26.4012
16.7013 | -34.9796 | -26.4088
16.7074 | -34.9801 | -26.4164
16.7135 | -34.9806 | -26.4241
16.7196 | -34.981 | -26.4317
16.7257 | -34.9815 | -26.4393
16.7318 | -34.982 | -26.4469
16.7379 | -34.9825 | -26.4545
16.744 | -34.983 | -26.462
16.7501 | -34.9835 | -26.4695
16.7562 | -34.9839 | -26.4772
16.7623 | -34.9844 | -26.4849
16.7684 | -34.9849 | -26.4924
16.7745 | -34.9854 | -26.5
16.7806 | -34.9858 | -26.5074
16.7867 | -34.9863 | -26.515
16.7928 | -34.9868 | -26.5228
16.7989 | -34.9873 | -26.5305
16.805 | -34.9878 | -26.5382
16.8111 | -34.9883 | -26.5457
16.8172 | -34.9888 | -26.5533
16.8233 | -34.9893 | -26.561
16.8294 | -34.9898 | -26.5686
16.8356 | -34.9903 | -26.5763
16.8417 | -34.9907 | -26.5839
16.8478 | -34.9912 | -26.5915
16.8539 | -34.9917 | -26.5991
16.86 | -34.9922 | -26.6068
16.8661 | -34.9927 | -26.6143
16.8722 | -34.9931 | -26.6221
16.8783 | -34.9936 | -26.6297
16.8844 | -34.9941 | -26.6372
16.8905 | -34.9946 | -26.6448
16.8966 | -34.995 | -26.6525
16.9027 | -34.9955 | -26.6601
16.9088 | -34.996 | -26.6678
16.9149 | -34.9965 | -26.6754
16.921 | -34.997 | -26.683
16.9271 | -34.9975 | -26.6906
16.9332 | -34.998 | -26.6982
16.9393 | -34.9985 | -26.706
16.9454 | -34.9989 | -26.7135
16.9515 | -34.9994 | -26.7212
16.9576 | -34.9999 | -26.7286
16.9637 | -35.0004 | -26.7362
16.9698 | -35.0009 | -26.7439
16.9759 | -35.0013 | -26.7517
16.9821 | -35.0018 | -26.7593
16.9882 | -35.0023 | -26.7667
16.9943 | -35.0028 | -26.7744
17.0004 | -35.0033 | -26.782
17.0065 | -35.0038 | -26.7895
17.0126 | -35.0043 | -26.797
17.0187 | -35.0047 | -26.8045
17.0248 | -35.0052 | -26.8118
17.0309 | -35.0057 | -26.8194
17.037 | -35.0061 | -26.827
17.0431 | -35.0066 | -26.8346
17.0492 | -35.0071 | -26.8423
17.0553 | -35.0077 | -26.8498
17.0614 | -35.0081 | -26.8573
17.0675 | -35.0086 | -26.8648
17.0736 | -35.0091 | -26.8724
17.0797 | -35.0096 | -26.88
17.0858 | -35.0101 | -26.8876
17.0919 | -35.0106 | -26.8951
17.098 | -35.011 | -26.9026
17.1041 | -35.0115 | -26.9102
17.1102 | -35.012 | -26.9177
17.1163 | -35.0125 | -26.9251
17.1225 | -35.013 | -26.9325
17.1286 | -35.0135 | -26.9397
17.1347 | -35.0139 | -26.9472
17.1408 | -35.0144 | -26.9548
17.1469 | -35.0149 | -26.9624
17.153 | -35.0154 | -26.9701
17.1591 | -35.0159 | -26.9777
17.1652 | -35.0164 | -26.9851
17.1713 | -35.0168 | -26.9926
17.1774 | -35.0173 | -27.0003
17.1835 | -35.0178 | -27.0078
17.1896 | -35.0182 | -27.0155
17.1957 | -35.0187 | -27.0229
17.2018 | -35.0192 | -27.0304
17.2079 | -35.0197 | -27.0378
17.214 | -35.0202 | -27.0453
17.2201 | -35.0207 | -27.0528
17.2262 | -35.0211 | -27.0604
17.2323 | -35.0216 | -27.0678
17.2384 | -35.0221 | -27.075
17.2445 | -35.0226 | -27.0826
17.2506 | -35.0231 | -27.0901
17.2567 | -35.0236 | -27.0977
17.2628 | -35.0241 | -27.1052
17.269 | -35.0246 | -27.1125
17.2751 | -35.025 | -27.1201
17.2812 | -35.0255 | -27.1276
17.2873 | -35.026 | -27.1352
17.2934 | -35.0265 | -27.1428
17.2995 | -35.027 | -27.1504
17.3056 | -35.0275 | -27.1577
17.3117 | -35.0279 | -27.1653
17.3178 | -35.0284 | -27.1729
17.3239 | -35.0289 | -27.1804
17.33 | -35.0293 | -27.188
17.3361 | -35.0298 | -27.1955
17.3422 | -35.0303 | -27.2028
17.3483 | -35.0308 | -27.2102
17.3544 | -35.0313 | -27.2176
17.3605 | -35.0318 | -27.2251
17.3666 | -35.0322 | -27.2326
17.3727 | -35.0327 | -27.2402
17.3788 | -35.0332 | -27.2478
17.3849 | -35.0336 | -27.2552
17.391 | -35.0341 | -27.2626
17.3971 | -35.0346 | -27.2701
17.4032 | -35.035 | -27.2776
17.4094 | -35.0355 | -27.2849
17.4155 | -35.036 | -27.2924
17.4216 | -35.0365 | -27.2999
17.4277 | -35.037 | -27.3073
17.4338 | -35.0374 | -27.3149
17.4399 | -35.0379 | -27.3223
17.446 | -35.0384 | -27.3297
17.4521 | -35.0389 | -27.3371
17.4582 | -35.0393 | -27.3444
17.4643 | -35.0398 | -27.3521
17.4704 | -35.0403 | -27.3596
17.4765 | -35.0408 | -27.367
17.4826 | -35.0413 | -27.3742
17.4887 | -35.0417 | -27.3817
17.4948 | -35.0422 | -27.389
17.5009 | -35.0427 | -27.3965
17.507 | -35.0432 | -27.4037
17.5131 | -35.0436 | -27.4113
17.5192 | -35.0441 | -27.4188
17.5253 | -35.0446 | -27.4262
17.5314 | -35.0451 | -27.4336
17.5375 | -35.0456 | -27.4409
17.5436 | -35.046 | -27.4484
17.5497 | -35.0465 | -27.4558
17.5559 | -35.047 | -27.4629
17.562 | -35.0475 | -27.4704
17.5681 | -35.048 | -27.4779
17.5742 | -35.0485 | -27.4852
17.5803 | -35.049 | -27.4926
17.5864 | -35.0494 | -27.4997
17.5925 | -35.0499 | -27.507
17.5986 | -35.0504 | -27.5144
17.6047 | -35.0509 | -27.5217
17.6108 | -35.0513 | -27.5292
17.6169 | -35.0518 | -27.5364
17.623 | -35.0524 | -27.5436
17.6291 | -35.0528 | -27.5509
17.6352 | -35.0533 | -27.5584
17.6413 | -35.0538 | -27.5657
17.6474 | -35.0543 | -27.5733
17.6535 | -35.0547 | -27.5809
17.6596 | -35.0553 | -27.5876
17.6657 | -35.0558 | -27.595
17.6718 | -35.0563 | -27.6025
17.6779 | -35.0567 | -27.61
17.684 | -35.0572 | -27.6175
17.6901 | -35.0577 | -27.6249
17.6963 | -35.0582 | -27.6321
17.7024 | -35.0586 | -27.6394
17.7085 | -35.0591 | -27.6468
17.7146 | -35.0596 | -27.6542
17.7207 | -35.0601 | -27.6614
17.7268 | -35.0605 | -27.6686
17.7329 | -35.061 | -27.676
17.739 | -35.0614 | -27.6833
17.7451 | -35.0619 | -27.6905
17.7512 | -35.0624 | -27.6978
17.7573 | -35.0629 | -27.7052
17.7634 | -35.0634 | -27.7122
17.7695 | -35.0638 | -27.7193
17.7756 | -35.0643 | -27.7265
17.7817 | -35.0648 | -27.7338
17.7878 | -35.0653 | -27.7409
17.7939 | -35.0657 | -27.7482
17.8 | -35.0662 | -27.7556
17.8061 | -35.0667 | -27.7629
17.8122 | -35.0671 | -27.7701
17.8183 | -35.0676 | -27.7774
17.8244 | -35.0681 | -27.7847
17.8305 | -35.0686 | -27.7921
17.8366 | -35.0691 | -27.7992
17.8428 | -35.0695 | -27.8065
17.8489 | -35.07 | -27.8138
17.855 | -35.0705 | -27.8213
17.8611 | -35.0709 | -27.8284
17.8672 | -35.0714 | -27.8355
17.8733 | -35.0719 | -27.8428
17.8794 | -35.0724 | -27.8501
17.8855 | -35.0728 | -27.8573
17.8916 | -35.0733 | -27.8646
17.8977 | -35.0738 | -27.8718
17.9038 | -35.0743 | -27.8789
17.9099 | -35.0748 | -27.8863
17.916 | -35.0753 | -27.8935
17.9221 | -35.0758 | -27.9008
17.9282 | -35.0762 | -27.9081
17.9343 | -35.0767 | -27.9154
17.9404 | -35.0771 | -27.9228
17.9465 | -35.0776 | -27.93
17.9526 | -35.0781 | -27.9373
17.9587 | -35.0786 | -27.9446
17.9648 | -35.0791 | -27.9518
17.9709 | -35.0795 | -27.9591
17.977 | -35.08 | -27.9662
17.9832 | -35.0805 | -27.9733
17.9893 | -35.081 | -27.9806
17.9954 | -35.0814 | -27.9882
18.0015 | -35.0819 | -27.9955
18.0076 | -35.0824 | -28.0025
18.0137 | -35.0829 | -28.0098
18.0198 | -35.0833 | -28.017
18.0259 | -35.0838 | -28.0244
18.032 | -35.0843 | -28.0317
18.0381 | -35.0848 | -28.0389
18.0442 | -35.0852 | -28.0461
18.0503 | -35.0857 | -28.0534
18.0564 | -35.0862 | -28.0606
18.0625 | -35.0867 | -28.0678
18.0686 | -35.0872 | -28.0751
18.0747 | -35.0876 | -28.0823
18.0808 | -35.0881 | -28.0892
18.0869 | -35.0886 | -28.0966
18.093 | -35.089 | -28.1038
18.0991 | -35.0895 | -28.1111
18.1052 | -35.09 | -28.1185
18.1113 | -35.0905 | -28.1257
18.1174 | -35.091 | -28.1328
18.1236 | -35.0914 | -28.14
18.1297 | -35.0919 | -28.1473
18.1358 | -35.0924 | -28.1545
18.1419 | -35.0928 | -28.1617
18.148 | -35.0933 | -28.1691
18.1541 | -35.0938 | -28.1758
18.1602 | -35.0942 | -28.1831
18.1663 | -35.0947 | -28.1904
18.1724 | -35.0952 | -28.1975
18.1785 | -35.0956 | -28.2047
18.1846 | -35.0961 | -28.2118
18.1907 | -35.0966 | -28.219
18.1968 | -35.0971 | -28.2262
18.2029 | -35.0975 | -28.2335
18.209 | -35.098 | -28.2407
18.2151 | -35.0985 | -28.2479
18.2212 | -35.099 | -28.2549
18.2273 | -35.0995 | -28.2621
18.2334 | -35.1 | -28.2691
18.2395 | -35.1005 | -28.2764
18.2456 | -35.1009 | -28.2838
18.2517 | -35.1014 | -28.2909
18.2578 | -35.1019 | -28.2978
18.2639 | -35.1024 | -28.305
18.2701 | -35.1029 | -28.3123
18.2762 | -35.1034 | -28.3194
18.2823 | -35.1038 | -28.3266
18.2884 | -35.1043 | -28.3337
18.2945 | -35.1047 | -28.3406
18.3006 | -35.1052 | -28.3477
18.3067 | -35.1057 | -28.3549
18.3128 | -35.1062 | -28.3621
18.3189 | -35.1066 | -28.3693
18.325 | -35.1071 | -28.3762
18.3311 | -35.1076 | -28.3834
18.3372 | -35.108 | -28.3906
18.3433 | -35.1085 | -28.3979
18.3494 | -35.109 | -28.4051
18.3555 | -35.1095 | -28.4122
18.3616 | -35.1099 | -28.4188
18.3677 | -35.1104 | -28.426
18.3738 | -35.1108 | -28.4332
18.3799 | -35.1113 | -28.4402
18.386 | -35.1118 | -28.4474
18.3921 | -35.1123 | -28.4547
18.3982 | -35.1127 | -28.4618
18.4043 | -35.1132 | -28.469
18.4105 | -35.1137 | -28.4761
18.4166 | -35.1142 | -28.4831
18.4227 | -35.1146 | -28.4902
18.4288 | -35.1151 | -28.4973
18.4349 | -35.1156 | -28.5042
18.441 | -35.1161 | -28.5114
18.4471 | -35.1165 | -28.5186
18.4532 | -35.117 | -28.5256
18.4593 | -35.1175 | -28.5327
18.4654 | -35.118 | -28.5395
18.4715 | -35.1184 | -28.5467
18.4776 | -35.1189 | -28.5539
18.4837 | -35.1194 | -28.5609
18.4898 | -35.1198 | -28.5683
18.4959 | -35.1203 | -28.5755
18.502 | -35.1208 | -28.5823
18.5081 | -35.1213 | -28.5893
18.5142 | -35.1217 | -28.5967
18.5203 | -35.1222 | -28.6035
18.5264 | -35.1226 | -28.6107
18.5325 | -35.1231 | -28.6178
18.5386 | -35.1236 | -28.6246
18.5447 | -35.1241 | -28.6317
18.5508 | -35.1245 | -28.6387
18.557 | -35.125 | -28.6458
18.5631 | -35.1255 | -28.6528
18.5692 | -35.1259 | -28.6599
18.5753 | -35.1264 | -28.6669
18.5814 | -35.1269 | -28.674
18.5875 | -35.1274 | -28.681
18.5936 | -35.1278 | -28.6881
18.5997 | -35.1283 | -28.6954
18.6058 | -35.1288 | -28.7023
18.6119 | -35.1293 | -28.7093
18.618 | -35.1297 | -28.7164
18.6241 | -35.1302 | -28.7236
18.6302 | -35.1307 | -28.7307
18.6363 | -35.1312 | -28.738
18.6424 | -35.1316 | -28.745
18.6485 | -35.1321 | -28.7518
18.6546 | -35.1325 | -28.7589
18.6607 | -35.133 | -28.7661
18.6668 | -35.1335 | -28.773
18.6729 | -35.134 | -28.7801
18.679 | -35.1345 | -28.7871
18.6851 | -35.135 | -28.7942
18.6912 | -35.1355 | -28.8011
18.6974 | -35.1359 | -28.8082
18.7035 | -35.1365 | -28.8151
18.7096 | -35.1369 | -28.8222
18.7157 | -35.1374 | -28.8291
18.7218 | -35.1378 | -28.836
18.7279 | -35.1383 | -28.8429
18.734 | -35.1388 | -28.8499
18.7401 | -35.1392 | -28.8569
18.7462 | -35.1397 | -28.8636
18.7523 | -35.1402 | -28.8709
18.7584 | -35.1407 | -28.8779
18.7645 | -35.1412 | -28.8852
18.7706 | -35.1417 | -28.8924
18.7767 | -35.1422 | -28.8994
18.7828 | -35.1426 | -28.9064
18.7889 | -35.1431 | -28.9133
18.795 | -35.1436 | -28.9204
18.8011 | -35.1441 | -28.9274
18.8072 | -35.1446 | -28.9345
18.8133 | -35.145 | -28.9413
18.8194 | -35.1455 | -28.9481
18.8255 | -35.146 | -28.9552
18.8316 | -35.1464 | -28.9622
18.8377 | -35.1469 | -28.9692
18.8439 | -35.1474 | -28.976
18.85 | -35.1478 | -28.9829
18.8561 | -35.1483 | -28.9898
18.8622 | -35.1488 | -28.9969
18.8683 | -35.1493 | -29.004
18.8744 | -35.1498 | -29.0111
18.8805 | -35.1503 | -29.0182
18.8866 | -35.1508 | -29.0252
18.8927 | -35.1512 | -29.0319
18.8988 | -35.1517 | -29.0386
18.9049 | -35.1522 | -29.0457
18.911 | -35.1527 | -29.0527
18.9171 | -35.1531 | -29.0598
18.9232 | -35.1536 | -29.0664
18.9293 | -35.1541 | -29.0732
18.9354 | -35.1545 | -29.0804
18.9415 | -35.155 | -29.0873
18.9476 | -35.1555 | -29.0944
18.9537 | -35.156 | -29.1014
18.9598 | -35.1564 | -29.1082
18.9659 | -35.1569 | -29.115
18.972 | -35.1573 | -29.1218
18.9781 | -35.1578 | -29.1288
18.9843 | -35.1583 | -29.1357
18.9904 | -35.1588 | -29.1427
18.9965 | -35.1592 | -29.1494
19.0026 | -35.1597 | -29.1564
19.0087 | -35.1602 | -29.1633
19.0148 | -35.1606 | -29.1701
19.0209 | -35.1611 | -29.177
19.027 | -35.1615 | -29.1841
19.0331 | -35.162 | -29.1907
19.0392 | -35.1625 | -29.1976
19.0453 | -35.1629 | -29.2046
19.0514 | -35.1634 | -29.2115
19.0575 | -35.1638 | -29.2186
19.0636 | -35.1643 | -29.2253
19.0697 | -35.1648 | -29.2322
19.0758 | -35.1653 | -29.2391
19.0819 | -35.1657 | -29.246
19.088 | -35.1662 | -29.2528
19.0941 | -35.1667 | -29.2599
19.1002 | -35.1672 | -29.2666
19.1063 | -35.1676 | -29.2736
19.1124 | -35.1681 | -29.2805
19.1185 | -35.1686 | -29.2874
19.1246 | -35.1691 | -29.2943
19.1308 | -35.1696 | -29.3012
19.1369 | -35.17 | -29.3079
19.143 | -35.1705 | -29.3148
19.1491 | -35.1709 | -29.3218
19.1552 | -35.1714 | -29.3288
19.1613 | -35.1719 | -29.3358
19.1674 | -35.1723 | -29.3429
19.1735 | -35.1728 | -29.3495
19.1796 | -35.1733 | -29.3565
19.1857 | -35.1737 | -29.3633
19.1918 | -35.1742 | -29.3702
19.1979 | -35.1746 | -29.3769
19.204 | -35.1751 | -29.3836
19.2101 | -35.1756 | -29.3906
19.2162 | -35.176 | -29.3975
19.2223 | -35.1765 | -29.4044
19.2284 | -35.1769 | -29.4113
19.2345 | -35.1774 | -29.4184
19.2406 | -35.1779 | -29.425
19.2467 | -35.1784 | -29.432
19.2528 | -35.1788 | -29.4393
19.2589 | -35.1793 | -29.4462
19.265 | -35.1798 | -29.4531
19.2712 | -35.1802 | -29.46
19.2773 | -35.1807 | -29.4664
19.2834 | -35.1812 | -29.4732
19.2895 | -35.1816 | -29.4801
19.2956 | -35.1821 | -29.4872
19.3017 | -35.1825 | -29.4941
19.3078 | -35.183 | -29.5011
19.3139 | -35.1835 | -29.5078
19.32 | -35.1839 | -29.5146
19.3261 | -35.1844 | -29.5215
19.3322 | -35.1849 | -29.528
19.3383 | -35.1853 | -29.5349
19.3444 | -35.1857 | -29.5415
19.3505 | -35.1862 | -29.5484
19.3566 | -35.1866 | -29.5554
19.3627 | -35.1871 | -29.5623
19.3688 | -35.1876 | -29.5694
19.3749 | -35.1881 | -29.5762
19.381 | -35.1885 | -29.5828
19.3871 | -35.189 | -29.5896
19.3932 | -35.1895 | -29.5964
19.3993 | -35.19 | -29.6036
19.4054 | -35.1904 | -29.6103
19.4115 | -35.1909 | -29.6173
19.4177 | -35.1913 | -29.6241
19.4238 | -35.1918 | -29.6309
19.4299 | -35.1922 | -29.638
19.436 | -35.1927 | -29.6448
19.4421 | -35.1932 | -29.6519
19.4482 | -35.1936 | -29.6588
19.4543 | -35.1941 | -29.6654
19.4604 | -35.1946 | -29.6725
19.4665 | -35.195 | -29.6795
19.4726 | -35.1955 | -29.6864
19.4787 | -35.1959 | -29.6933
19.4848 | -35.1964 | -29.7001
19.4909 | -35.1969 | -29.7068
19.497 | -35.1974 | -29.7137
19.5031 | -35.1979 | -29.7205
19.5092 | -35.1983 | -29.7274
19.5153 | -35.1988 | -29.7342
19.5214 | -35.1992 | -29.7409
19.5275 | -35.1997 | -29.7478
19.5336 | -35.2002 | -29.7548
19.5397 | -35.2007 | -29.7617
19.5458 | -35.2011 | -29.7685
19.5519 | -35.2015 | -29.7753
19.5581 | -35.202 | -29.782
19.5642 | -35.2025 | -29.789
19.5703 | -35.203 | -29.7959
19.5764 | -35.2034 | -29.8026
19.5825 | -35.2039 | -29.8093
19.5886 | -35.2043 | -29.8163
19.5947 | -35.2048 | -29.8231
19.6008 | -35.2052 | -29.83
19.6069 | -35.2057 | -29.8368
19.613 | -35.2062 | -29.8435
19.6191 | -35.2066 | -29.8504
19.6252 | -35.2071 | -29.8571
19.6313 | -35.2076 | -29.8638
19.6374 | -35.208 | -29.8708
19.6435 | -35.2085 | -29.8777
19.6496 | -35.2089 | -29.8842
19.6557 | -35.2094 | -29.8912
19.6618 | -35.2098 | -29.8979
19.6679 | -35.2103 | -29.9048
19.674 | -35.2107 | -29.9116
19.6801 | -35.2112 | -29.9183
19.6862 | -35.2116 | -29.9251
19.6923 | -35.2121 | -29.9321
19.6984 | -35.2126 | -29.939
19.7046 | -35.213 | -29.9458
19.7107 | -35.2135 | -29.9526
19.7168 | -35.214 | -29.9595
19.7229 | -35.2145 | -29.9661
19.729 | -35.2149 | -29.9729
19.7351 | -35.2154 | -29.9794
19.7412 | -35.2159 | -29.9863
19.7473 | -35.2164 | -29.9931
19.7534 | -35.2168 | -30.0001
19.7595 | -35.2173 | -30.0071
19.7656 | -35.2178 | -30.0138
19.7717 | -35.2182 | -30.0205
19.7778 | -35.2187 | -30.0275
19.7839 | -35.2192 | -30.0343
19.79 | -35.2196 | -30.041
19.7961 | -35.2201 | -30.0478
19.8022 | -35.2205 | -30.0547
19.8083 | -35.221 | -30.0612
19.8144 | -35.2215 | -30.068
19.8205 | -35.2219 | -30.0748
19.8266 | -35.2224 | -30.0815
19.8327 | -35.2229 | -30.0886
19.8388 | -35.2233 | -30.095
19.845 | -35.2238 | -30.1017
19.8511 | -35.2243 | -30.1085
19.8572 | -35.2247 | -30.1154
19.8633 | -35.2252 | -30.1222
19.8694 | -35.2256 | -30.1293
19.8755 | -35.2261 | -30.1359
19.8816 | -35.2266 | -30.1427
19.8877 | -35.2271 | -30.1494
19.8938 | -35.2275 | -30.1563
19.8999 | -35.228 | -30.1631
19.906 | -35.2285 | -30.17
19.9121 | -35.2289 | -30.1767
19.9182 | -35.2294 | -30.1836
19.9243 | -35.2299 | -30.1904
19.9304 | -35.2303 | -30.197
19.9365 | -35.2308 | -30.204
19.9426 | -35.2313 | -30.2107
19.9487 | -35.2317 | -30.2175
19.9548 | -35.2322 | -30.2242
19.9609 | -35.2326 | -30.2309
19.967 | -35.2331 | -30.2376
19.9731 | -35.2335 | -30.2443
19.9792 | -35.234 | -30.2511
19.9853 | -35.2344 | -30.2577
19.9915 | -35.2349 | -30.2643
19.9976 | -35.2353 | -30.2712
20.0037 | -35.2358 | -30.2781
20.0098 | -35.2362 | -30.2848
20.0159 | -35.2367 | -30.2917
20.022 | -35.2372 | -30.2985
20.0281 | -35.2376 | -30.3056
20.0342 | -35.2381 | -30.3123
20.0403 | -35.2385 | -30.3192
20.0464 | -35.239 | -30.326
20.0525 | -35.2395 | -30.3321
20.0586 | -35.24 | -30.339
20.0647 | -35.2404 | -30.3459
20.0708 | -35.2409 | -30.3527
20.0769 | -35.2413 | -30.3594
20.083 | -35.2418 | -30.3661
20.0891 | -35.2423 | -30.373
20.0952 | -35.2427 | -30.3796
20.1013 | -35.2432 | -30.3867
20.1074 | -35.2437 | -30.3932
20.1135 | -35.2441 | -30.3999
20.1196 | -35.2446 | -30.4067
20.1257 | -35.2451 | -30.4135
20.1319 | -35.2455 | -30.4203
20.138 | -35.246 | -30.4269
20.1441 | -35.2464 | -30.4337
20.1502 | -35.2469 | -30.4405
20.1563 | -35.2473 | -30.4474
20.1624 | -35.2478 | -30.4543
20.1685 | -35.2482 | -30.461
20.1746 | -35.2487 | -30.4678
20.1807 | -35.2491 | -30.4744
20.1868 | -35.2496 | -30.4813
20.1929 | -35.2501 | -30.4877
20.199 | -35.2506 | -30.4945
20.2051 | -35.2511 | -30.5013
20.2112 | -35.2516 | -30.5082
20.2173 | -35.252 | -30.5151
20.2234 | -35.2525 | -30.522
20.2295 | -35.253 | -30.5284
20.2356 | -35.2535 | -30.5353
20.2417 | -35.2539 | -30.542
20.2478 | -35.2544 | -30.5485
20.2539 | -35.2549 | -30.5551
20.26 | -35.2553 | -30.5617
20.2661 | -35.2558 | -30.5684
20.2723 | -35.2562 | -30.5751
20.2784 | -35.2566 | -30.5819
20.2845 | -35.2571 | -30.5885
20.2906 | -35.2575 | -30.5951
20.2967 | -35.258 | -30.6018
20.3028 | -35.2584 | -30.6085
20.3089 | -35.2589 | -30.6152
20.315 | -35.2593 | -30.622
20.3211 | -35.2598 | -30.6287
20.3272 | -35.2603 | -30.6353
20.3333 | -35.2607 | -30.6421
20.3394 | -35.2611 | -30.6489
20.3455 | -35.2616 | -30.6555
20.3516 | -35.262 | -30.6621
20.3577 | -35.2624 | -30.6689
20.3638 | -35.2629 | -30.6754
20.3699 | -35.2634 | -30.682
20.376 | -35.2638 | -30.6886
20.3821 | -35.2643 | -30.6953
20.3882 | -35.2647 | -30.7018
20.3943 | -35.2652 | -30.7085
20.4004 | -35.2656 | -30.7152
20.4065 | -35.266 | -30.722
20.4126 | -35.2665 | -30.7288
20.4188 | -35.267 | -30.7355
20.4249 | -35.2674 | -30.7423
20.431 | -35.2678 | -30.749
20.4371 | -35.2682 | -30.7553
20.4432 | -35.2687 | -30.7623
20.4493 | -35.2691 | -30.7691
20.4554 | -35.2696 | -30.7756
20.4615 | -35.27 | -30.7821
20.4676 | -35.2705 | -30.7891
20.4737 | -35.2709 | -30.7959
20.4798 | -35.2713 | -30.8025
20.4859 | -35.2718 | -30.8092
20.492 | -35.2723 | -30.8158
20.4981 | -35.2727 | -30.8224
20.5042 | -35.2732 | -30.8292
20.5103 | -35.2736 | -30.8359
20.5164 | -35.274 | -30.8428
20.5225 | -35.2745 | -30.8495
20.5286 | -35.2749 | -30.8565
20.5347 | -35.2754 | -30.8632
20.5408 | -35.2758 | -30.8698
20.5469 | -35.2762 | -30.8765
20.553 | -35.2767 | -30.8833
20.5592 | -35.2772 | -30.8898
20.5653 | -35.2777 | -30.8965
20.5714 | -35.2781 | -30.9037
20.5775 | -35.2786 | -30.9101
20.5836 | -35.279 | -30.9166
20.5897 | -35.2795 | -30.9234
20.5958 | -35.2799 | -30.9302
20.6019 | -35.2804 | -30.9368
20.608 | -35.2808 | -30.9438
20.6141 | -35.2813 | -30.9507
20.6202 | -35.2818 | -30.9574
20.6263 | -35.2822 | -30.9639
20.6324 | -35.2826 | -30.9707
20.6385 | -35.2831 | -30.9775
20.6446 | -35.2835 | -30.9841
20.6507 | -35.284 | -30.9906
20.6568 | -35.2845 | -30.9975
20.6629 | -35.2849 | -31.0042
20.669 | -35.2854 | -31.0111
20.6751 | -35.2858 | -31.0178
20.6812 | -35.2863 | -31.0244
20.6873 | -35.2867 | -31.0311
20.6934 | -35.2871 | -31.0378
20.6995 | -35.2876 | -31.0443
20.7057 | -35.2881 | -31.0511
20.7118 | -35.2886 | -31.0577
20.7179 | -35.289 | -31.0641
20.724 | -35.2895 | -31.071
20.7301 | -35.2899 | -31.0777
20.7362 | -35.2903 | -31.0843
20.7423 | -35.2908 | -31.0909
20.7484 | -35.2913 | -31.0975
20.7545 | -35.2917 | -31.1041
20.7606 | -35.2922 | -31.1107
20.7667 | -35.2926 | -31.1174
20.7728 | -35.2931 | -31.124
20.7789 | -35.2935 | -31.1308
20.785 | -35.294 | -31.1375
20.7911 | -35.2945 | -31.144
20.7972 | -35.2949 | -31.1507
20.8033 | -35.2954 | -31.1576
20.8094 | -35.2958 | -31.1642
20.8155 | -35.2962 | -31.1707
20.8216 | -35.2967 | -31.1774
20.8277 | -35.2971 | -31.1841
20.8338 | -35.2976 | -31.1907
20.8399 | -35.2981 | -31.1972
20.8461 | -35.2985 | -31.204
20.8522 | -35.299 | -31.2105
20.8583 | -35.2995 | -31.2172
20.8644 | -35.2999 | -31.2242
20.8705 | -35.3004 | -31.2311
20.8766 | -35.3009 | -31.2382
20.8827 | -35.3013 | -31.2446
20.8888 | -35.3018 | -31.2513
20.8949 | -35.3022 | -31.2578
20.901 | -35.3027 | -31.2648
20.9071 | -35.3032 | -31.2714
20.9132 | -35.3037 | -31.2779
20.9193 | -35.3041 | -31.2846
20.9254 | -35.3046 | -31.2912
20.9315 | -35.3051 | -31.2977
20.9376 | -35.3055 | -31.3044
20.9437 | -35.3059 | -31.3112
20.9498 | -35.3064 | -31.3179
20.9559 | -35.3069 | -31.3248
20.962 | -35.3073 | -31.3308
20.9681 | -35.3078 | -31.3372
20.9742 | -35.3083 | -31.3439
20.9803 | -35.3087 | -31.3504
20.9864 | -35.3092 | -31.3573
20.9926 | -35.3096 | -31.3638
20.9987 | -35.31 | -31.371
21.0048 | -35.3105 | -31.3778
21.0109 | -35.3109 | -31.3842
21.017 | -35.3114 | -31.3908
21.0231 | -35.3119 | -31.3973
21.0292 | -35.3123 | -31.4043
21.0353 | -35.3128 | -31.4113
21.0414 | -35.3132 | -31.4185
21.0475 | -35.3136 | -31.425
21.0536 | -35.314 | -31.4316
21.0597 | -35.3145 | -31.4378
21.0658 | -35.3149 | -31.4442
21.0719 | -35.3154 | -31.4505
21.078 | -35.3159 | -31.4576
21.0841 | -35.3162 | -31.4636
21.0902 | -35.3166 | -31.4701
21.0963 | -35.3171 | -31.4765
21.1024 | -35.3177 | -31.4824
21.1085 | -35.3182 | -31.4885
21.1146 | -35.3185 | -31.4955
21.1207 | -35.319 | -31.5024
21.1268 | -35.3195 | -31.5095
21.133 | -35.32 | -31.5163
21.1391 | -35.3205 | -31.5227
21.1452 | -35.321 | -31.5294
21.1513 | -35.3214 | -31.5356
21.1574 | -35.3218 | -31.542
21.1635 | -35.3223 | -31.5488
21.1696 | -35.3227 | -31.5555
21.1757 | -35.3232 | -31.5617
21.1818 | -35.3236 | -31.5684
21.1879 | -35.324 | -31.5753
21.194 | -35.3245 | -31.5819
21.2001 | -35.3249 | -31.588
21.2062 | -35.3254 | -31.5951
21.2123 | -35.3256 | -31.6024
21.2184 | -35.3263 | -31.6082
21.2245 | -35.3272 | -31.6173
21.2306 | -35.3275 | -31.6284
21.2367 | -35.3278 | -31.6346
21.2428 | -35.3282 | -31.641
21.2489 | -35.3286 | -31.6478
21.255 | -35.3291 | -31.6544
21.2611 | -35.3295 | -31.6608
21.2672 | -35.33 | -31.6673
21.2733 | -35.3304 | -31.674
21.2795 | -35.3309 | -31.6802
21.2856 | -35.3314 | -31.6873
21.2917 | -35.3319 | -31.6943
21.2978 | -35.3323 | -31.701
21.3039 | -35.3328 | -31.7078
21.31 | -35.3332 | -31.7145
21.3161 | -35.3337 | -31.7209
21.3222 | -35.3341 | -31.7278
21.3283 | -35.3345 | -31.7344
21.3344 | -35.335 | -31.7412
21.3405 | -35.3354 | -31.7478
21.3466 | -35.3358 | -31.7543
21.3527 | -35.3363 | -31.7606
21.3588 | -35.3368 | -31.7674
21.3649 | -35.3372 | -31.774
21.371 | -35.3377 | -31.7806
21.3771 | -35.3381 | -31.7873
21.3832 | -35.3386 | -31.7939
21.3893 | -35.3391 | -31.8008
21.3954 | -35.3396 | -31.8076
21.4015 | -35.3401 | -31.8145
21.4076 | -35.3405 | -31.8212
21.4137 | -35.3409 | -31.8276
21.4199 | -35.3414 | -31.8343
21.426 | -35.3418 | -31.8404
21.4321 | -35.3422 | -31.8473
21.4382 | -35.3426 | -31.8542
21.4443 | -35.3431 | -31.8611
21.4504 | -35.3436 | -31.8678
21.4565 | -35.3441 | -31.8744
21.4626 | -35.3445 | -31.8812
21.4687 | -35.3449 | -31.8878
21.4748 | -35.3454 | -31.8944
21.4809 | -35.3459 | -31.901
21.487 | -35.3463 | -31.9074
21.4931 | -35.3468 | -31.9138
21.4992 | -35.3472 | -31.9204
21.5053 | -35.3477 | -31.927
21.5114 | -35.3481 | -31.9335
21.5175 | -35.3486 | -31.9402
21.5236 | -35.349 | -31.9468
21.5297 | -35.3494 | -31.9533
21.5358 | -35.3498 | -31.96
21.5419 | -35.3503 | -31.9667
21.548 | -35.3508 | -31.9734
21.5541 | -35.3512 | -31.9801
21.5602 | -35.3517 | -31.9868
21.5664 | -35.3521 | -31.9933
21.5725 | -35.3526 | -32.0001
21.5786 | -35.3531 | -32.0067
21.5847 | -35.3535 | -32.0135
21.5908 | -35.354 | -32.0202
21.5969 | -35.3545 | -32.0268
21.603 | -35.355 | -32.0334
21.6091 | -35.3555 | -32.0402
21.6152 | -35.3559 | -32.0465
21.6213 | -35.3564 | -32.0536
21.6274 | -35.3568 | -32.06
21.6335 | -35.3572 | -32.0662
21.6396 | -35.3577 | -32.0726
21.6457 | -35.3581 | -32.0793
21.6518 | -35.3585 | -32.0859
21.6579 | -35.359 | -32.0923
21.664 | -35.3594 | -32.0993
21.6701 | -35.3599 | -32.1055
21.6762 | -35.3603 | -32.1122
21.6823 | -35.3607 | -32.119
21.6884 | -35.3612 | -32.1259
21.6945 | -35.3616 | -32.1324
21.7006 | -35.362 | -32.1391
21.7068 | -35.3625 | -32.1457
21.7129 | -35.363 | -32.1524
21.719 | -35.3635 | -32.1594
21.7251 | -35.3639 | -32.1661
21.7312 | -35.3644 | -32.1729
21.7373 | -35.3648 | -32.1795
21.7434 | -35.3652 | -32.1863
21.7495 | -35.3656 | -32.1927
21.7556 | -35.3661 | -32.1996
21.7617 | -35.3665 | -32.2063
21.7678 | -35.367 | -32.2129
21.7739 | -35.3674 | -32.2194
21.78 | -35.3679 | -32.2259
21.7861 | -35.3683 | -32.2325
21.7922 | -35.3688 | -32.2392
21.7983 | -35.3692 | -32.2458
21.8044 | -35.3697 | -32.2523
21.8105 | -35.3701 | -32.2589
21.8166 | -35.3706 | -32.2652
21.8227 | -35.371 | -32.2718
21.8288 | -35.3715 | -32.2784
21.8349 | -35.3719 | -32.2854
21.841 | -35.3723 | -32.2921
21.8471 | -35.3728 | -32.2984
21.8533 | -35.3732 | -32.3049
21.8594 | -35.3737 | -32.3116
21.8655 | -35.3741 | -32.3181
21.8716 | -35.3746 | -32.3244
21.8777 | -35.3751 | -32.331
21.8838 | -35.3754 | -32.3375
21.8899 | -35.3759 | -32.3441
21.896 | -35.3764 | -32.3503
21.9021 | -35.3768 | -32.3575
21.9082 | -35.3772 | -32.3643
21.9143 | -35.3775 | -32.3708
21.9204 | -35.378 | -32.3774
21.9265 | -35.3784 | -32.3837
21.9326 | -35.3789 | -32.3903
21.9387 | -35.3793 | -32.3972
21.9448 | -35.3797 | -32.4041
21.9509 | -35.3802 | -32.4104
21.957 | -35.3806 | -32.4173
21.9631 | -35.3811 | -32.4239
21.9692 | -35.3816 | -32.4307
21.9753 | -35.382 | -32.4374
21.9814 | -35.3825 | -32.4444
21.9875 | -35.3829 | -32.4507
21.9937 | -35.3834 | -32.4575
21.9998 | -35.3838 | -32.4637
22.0059 | -35.3842 | -32.4704
22.012 | -35.3846 | -32.4768
22.0181 | -35.385 | -32.4834
22.0242 | -35.3855 | -32.4897
22.0303 | -35.3859 | -32.4962
22.0364 | -35.3864 | -32.5029
22.0425 | -35.3868 | -32.5094
22.0486 | -35.3872 | -32.516
22.0547 | -35.3877 | -32.5225
22.0608 | -35.3881 | -32.5291
22.0669 | -35.3885 | -32.5355
22.073 | -35.3889 | -32.542
22.0791 | -35.3894 | -32.5485
22.0852 | -35.3899 | -32.555
22.0913 | -35.3904 | -32.5615
22.0974 | -35.3908 | -32.5679
22.1035 | -35.3912 | -32.5743
22.1096 | -35.3917 | -32.581
22.1157 | -35.3921 | -32.5875
22.1218 | -35.3926 | -32.5942
22.1279 | -35.393 | -32.6004
22.134 | -35.3934 | -32.6071
22.1402 | -35.3939 | -32.6136
22.1463 | -35.3943 | -32.6205
22.1524 | -35.3947 | -32.6272
22.1585 | -35.3952 | -32.6338
22.1646 | -35.3956 | -32.6403
22.1707 | -35.3961 | -32.6472
22.1768 | -35.3966 | -32.6539
22.1829 | -35.397 | -32.6605
22.189 | -35.3974 | -32.6671
22.1951 | -35.3979 | -32.6739
22.2012 | -35.3983 | -32.6806
22.2073 | -35.3988 | -32.6873
22.2134 | -35.3992 | -32.6936
22.2195 | -35.3996 | -32.7001
22.2256 | -35.4001 | -32.7068
22.2317 | -35.4005 | -32.713
22.2378 | -35.401 | -32.7197
22.2439 | -35.4015 | -32.7261
22.25 | -35.4019 | -32.7326
22.2561 | -35.4023 | -32.7393
22.2622 | -35.4028 | -32.7462
22.2683 | -35.4033 | -32.7527
22.2744 | -35.4037 | -32.7592
22.2806 | -35.4042 | -32.7656
22.2867 | -35.4046 | -32.7718
22.2928 | -35.4051 | -32.7782
22.2989 | -35.4055 | -32.7849
22.305 | -35.406 | -32.7911
22.3111 | -35.4065 | -32.7975
22.3172 | -35.4069 | -32.8041
22.3233 | -35.4073 | -32.8105
22.3294 | -35.4078 | -32.8168
22.3355 | -35.4082 | -32.8234
22.3416 | -35.4087 | -32.8299
22.3477 | -35.4091 | -32.8365
22.3538 | -35.4095 | -32.8431
22.3599 | -35.4099 | -32.8497
22.366 | -35.4104 | -32.8563
22.3721 | -35.4108 | -32.8627
22.3782 | -35.4112 | -32.8691
22.3843 | -35.4117 | -32.8753
22.3904 | -35.4121 | -32.8819
22.3965 | -35.4125 | -32.8884
22.4026 | -35.413 | -32.8951
22.4087 | -35.4134 | -32.9014
22.4148 | -35.4138 | -32.9081
22.4209 | -35.4142 | -32.9146
22.4271 | -35.4146 | -32.921
22.4332 | -35.4151 | -32.9275
22.4393 | -35.4156 | -32.9341
22.4454 | -35.416 | -32.9403
22.4515 | -35.4164 | -32.947
22.4576 | -35.4169 | -32.9536
22.4637 | -35.4173 | -32.9601
22.4698 | -35.4177 | -32.9666
22.4759 | -35.4181 | -32.9732
22.482 | -35.4185 | -32.9799
22.4881 | -35.419 | -32.9865
22.4942 | -35.4194 | -32.9931
22.5003 | -35.4198 | -32.9998
22.5064 | -35.4203 | -33.006
22.5125 | -35.4207 | -33.0121
22.5186 | -35.4211 | -33.0187
22.5247 | -35.4216 | -33.0252
22.5308 | -35.422 | -33.0316
22.5369 | -35.4225 | -33.0383
22.543 | -35.4229 | -33.045
22.5491 | -35.4232 | -33.0508
22.5552 | -35.4237 | -33.0574
22.5613 | -35.4241 | -33.0638
22.5675 | -35.4246 | -33.0701
22.5736 | -35.425 | -33.0766
22.5797 | -35.4255 | -33.0832
22.5858 | -35.4259 | -33.0891
22.5919 | -35.4264 | -33.0953
22.598 | -35.4268 | -33.1018
22.6041 | -35.4273 | -33.1082
22.6102 | -35.4277 | -33.1146
22.6163 | -35.4282 | -33.121
22.6224 | -35.4286 | -33.1275
22.6285 | -35.429 | -33.134
22.6346 | -35.4295 | -33.1408
22.6407 | -35.4299 | -33.1472
22.6468 | -35.4303 | -33.1535
22.6529 | -35.4307 | -33.1596
22.659 | -35.4311 | -33.1662
22.6651 | -35.4315 | -33.1723
22.6712 | -35.432 | -33.179
22.6773 | -35.4324 | -33.1855
22.6834 | -35.4328 | -33.1918
22.6895 | -35.4332 | -33.1979
22.6956 | -35.4337 | -33.2045
22.7017 | -35.4342 | -33.2111
22.7079 | -35.4346 | -33.2175
22.714 | -35.435 | -33.2238
22.7201 | -35.4354 | -33.2302
22.7262 | -35.4359 | -33.2366
22.7323 | -35.4363 | -33.2431
22.7384 | -35.4367 | -33.2494
22.7445 | -35.4372 | -33.2561
22.7506 | -35.4376 | -33.2626
22.7567 | -35.4381 | -33.269
22.7628 | -35.4385 | -33.2754
22.7689 | -35.4389 | -33.282
22.775 | -35.4394 | -33.2882
22.7811 | -35.4397 | -33.2945
22.7872 | -35.4402 | -33.3011
22.7933 | -35.4406 | -33.3073
22.7994 | -35.441 | -33.3137
22.8055 | -35.4414 | -33.32
22.8116 | -35.4419 | -33.3264
22.8177 | -35.4423 | -33.3328
22.8238 | -35.4428 | -33.3392
22.8299 | -35.4433 | -33.3453
22.836 | -35.4437 | -33.3517
22.8421 | -35.4441 | -33.3583
22.8482 | -35.4446 | -33.3646
22.8544 | -35.445 | -33.3711
22.8605 | -35.4455 | -33.3774
22.8666 | -35.4458 | -33.3836
22.8727 | -35.4462 | -33.3901
22.8788 | -35.4466 | -33.3963
22.8849 | -35.447 | -33.4027
22.891 | -35.4475 | -33.4091
22.8971 | -35.4479 | -33.4153
22.9032 | -35.4484 | -33.4215
22.9093 | -35.4488 | -33.4277
22.9154 | -35.4493 | -33.4339
22.9215 | -35.4497 | -33.4406
22.9276 | -35.4501 | -33.4468
22.9337 | -35.4506 | -33.4526
22.9398 | -35.451 | -33.4593
22.9459 | -35.4515 | -33.4658
22.952 | -35.4519 | -33.4722
22.9581 | -35.4523 | -33.4785
22.9642 | -35.4527 | -33.4848
22.9703 | -35.4531 | -33.4906
22.9764 | -35.4536 | -33.4971
22.9825 | -35.454 | -33.5035
22.9886 | -35.4544 | -33.5098
22.9948 | -35.4548 | -33.5163
23.0009 | -35.4552 | -33.5229
23.007 | -35.4557 | -33.5291
23.0131 | -35.4561 | -33.5355
23.0192 | -35.4565 | -33.5419
23.0253 | -35.457 | -33.5483
23.0314 | -35.4574 | -33.5548
23.0375 | -35.4578 | -33.5611
23.0436 | -35.4582 | -33.5671
23.0497 | -35.4587 | -33.5735
23.0558 | -35.4591 | -33.5799
23.0619 | -35.4596 | -33.5861
23.068 | -35.46 | -33.5923
23.0741 | -35.4605 | -33.5983
23.0802 | -35.4609 | -33.6045
23.0863 | -35.4613 | -33.6109
23.0924 | -35.4617 | -33.617
23.0985 | -35.4621 | -33.6232
23.1046 | -35.4625 | -33.6297
23.1107 | -35.4629 | -33.6356
23.1168 | -35.4634 | -33.642
23.1229 | -35.4638 | -33.6482
23.129 | -35.4642 | -33.6542
23.1351 | -35.4646 | -33.6603
23.1413 | -35.465 | -33.6668
23.1474 | -35.4655 | -33.673
23.1535 | -35.4659 | -33.6794
23.1596 | -35.4663 | -33.6858
23.1657 | -35.4667 | -33.6921
23.1718 | -35.4671 | -33.6984
23.1779 | -35.4676 | -33.7048
23.184 | -35.4679 | -33.7111
23.1901 | -35.4683 | -33.7174
23.1962 | -35.4687 | -33.7239
23.2023 | -35.4692 | -33.7304
23.2084 | -35.4696 | -33.7366
23.2145 | -35.47 | -33.7424
23.2206 | -35.4704 | -33.7487
23.2267 | -35.4708 | -33.7551
23.2328 | -35.4713 | -33.7612
23.2389 | -35.4717 | -33.7673
23.245 | -35.4722 | -33.7734
23.2511 | -35.4726 | -33.7792
23.2572 | -35.473 | -33.7855
23.2633 | -35.4734 | -33.7917
23.2694 | -35.4738 | -33.7982
23.2755 | -35.4742 | -33.8046
23.2817 | -35.4747 | -33.8111
23.2878 | -35.4752 | -33.8172
23.2939 | -35.4756 | -33.8235
23.3 | -35.476 | -33.8297
23.3061 | -35.4764 | -33.836
23.3122 | -35.4768 | -33.8421
23.3183 | -35.4772 | -33.8485
23.3244 | -35.4777 | -33.8545
23.3305 | -35.4781 | -33.8609
23.3366 | -35.4785 | -33.8673
23.3427 | -35.4789 | -33.8734
23.3488 | -35.4794 | -33.8796
23.3549 | -35.4798 | -33.886
23.361 | -35.4801 | -33.892
23.3671 | -35.4806 | -33.8982
23.3732 | -35.4809 | -33.9044
23.3793 | -35.4814 | -33.9109
23.3854 | -35.4818 | -33.9172
23.3915 | -35.4822 | -33.9231
23.3976 | -35.4826 | -33.9296
23.4037 | -35.4831 | -33.936
23.4098 | -35.4835 | -33.942
23.4159 | -35.4839 | -33.9483
23.422 | -35.4843 | -33.9544
23.4282 | -35.4846 | -33.9604
23.4343 | -35.4851 | -33.9667
23.4404 | -35.4855 | -33.9731
23.4465 | -35.486 | -33.9795
23.4526 | -35.4864 | -33.9858
23.4587 | -35.4867 | -33.9919
23.4648 | -35.4872 | -33.998
23.4709 | -35.4876 | -34.0043
23.477 | -35.488 | -34.0106
23.4831 | -35.4885 | -34.017
23.4892 | -35.4888 | -34.0229
23.4953 | -35.4892 | -34.0292
23.5014 | -35.4896 | -34.0351
23.5075 | -35.49 | -34.0413
23.5136 | -35.4905 | -34.0472
23.5197 | -35.4909 | -34.0533
23.5258 | -35.4913 | -34.0597
23.5319 | -35.4917 | -34.0657
23.538 | -35.4921 | -34.0719
23.5441 | -35.4924 | -34.078
23.5502 | -35.4929 | -34.0842
23.5563 | -35.4933 | -34.0902
23.5624 | -35.4937 | -34.0965
23.5686 | -35.4941 | -34.1024
23.5747 | -35.4945 | -34.1084
23.5808 | -35.4949 | -34.1146
23.5869 | -35.4953 | -34.1207
23.593 | -35.4957 | -34.127
23.5991 | -35.4962 | -34.1331
23.6052 | -35.4966 | -34.1389
23.6113 | -35.497 | -34.145
23.6174 | -35.4974 | -34.1513
23.6235 | -35.4979 | -34.1575
23.6296 | -35.4983 | -34.1637
23.6357 | -35.4987 | -34.1698
23.6418 | -35.4991 | -34.1755
23.6479 | -35.4995 | -34.1816
23.654 | -35.5 | -34.1877
23.6601 | -35.5004 | -34.1937
23.6662 | -35.5007 | -34.1998
23.6723 | -35.5011 | -34.2055
23.6784 | -35.5015 | -34.2116
23.6845 | -35.5019 | -34.2177
23.6906 | -35.5024 | -34.2238
23.6967 | -35.5028 | -34.23
23.7028 | -35.5031 | -34.2358
23.7089 | -35.5035 | -34.2418
23.7151 | -35.5039 | -34.2479
23.7212 | -35.5043 | -34.254
23.7273 | -35.5048 | -34.2603
23.7334 | -35.5052 | -34.266
23.7395 | -35.5056 | -34.2719
23.7456 | -35.5061 | -34.2779
23.7517 | -35.5065 | -34.2843
23.7578 | -35.5069 | -34.2904
23.7639 | -35.5073 | -34.2964
23.77 | -35.5077 | -34.3025
23.7761 | -35.5081 | -34.3088
23.7822 | -35.5085 | -34.3146
23.7883 | -35.5089 | -34.3202
23.7944 | -35.5093 | -34.3262
23.8005 | -35.5096 | -34.3324
23.8066 | -35.5101 | -34.3384
23.8127 | -35.5105 | -34.3444
23.8188 | -35.5109 | -34.3506
23.8249 | -35.5113 | -34.3567
23.831 | -35.5117 | -34.3629
23.8371 | -35.5122 | -34.3689
23.8432 | -35.5126 | -34.3751
23.8493 | -35.513 | -34.3808
23.8555 | -35.5134 | -34.3871
23.8616 | -35.5138 | -34.3927
23.8677 | -35.5142 | -34.3987
23.8738 | -35.5147 | -34.405
23.8799 | -35.5151 | -34.4108
23.886 | -35.5154 | -34.417
23.8921 | -35.5159 | -34.4234
23.8982 | -35.5163 | -34.4292
23.9043 | -35.5167 | -34.4352
23.9104 | -35.5171 | -34.441
23.9165 | -35.5175 | -34.4471
23.9226 | -35.5179 | -34.4528
23.9287 | -35.5184 | -34.459
23.9348 | -35.5188 | -34.4651
23.9409 | -35.5191 | -34.4713
23.947 | -35.5195 | -34.4775
23.9531 | -35.5199 | -34.4829
23.9592 | -35.5203 | -34.4889
23.9653 | -35.5207 | -34.495
23.9714 | -35.5211 | -34.5011
23.9775 | -35.5216 | -34.507
23.9836 | -35.522 | -34.5126
23.9897 | -35.5224 | -34.5184
23.9958 | -35.5228 | -34.5246
24.002 | -35.5232 | -34.5306
24.0081 | -35.5236 | -34.5365
24.0142 | -35.524 | -34.5425
24.0203 | -35.5244 | -34.5484
24.0264 | -35.5248 | -34.5542
24.0325 | -35.5252 | -34.56
24.0386 | -35.5257 | -34.5659
24.0447 | -35.5261 | -34.5718
24.0508 | -35.5265 | -34.5778
24.0569 | -35.5269 | -34.5834
24.063 | -35.5273 | -34.589
24.0691 | -35.5278 | -34.5948
24.0752 | -35.5282 | -34.6009
24.0813 | -35.5286 | -34.6071
24.0874 | -35.529 | -34.6131
24.0935 | -35.5294 | -34.6192
24.0996 | -35.5297 | -34.6248
24.1057 | -35.5301 | -34.6308
24.1118 | -35.5305 | -34.6366
24.1179 | -35.5309 | -34.6424
24.124 | -35.5313 | -34.6483
24.1301 | -35.5318 | -34.6539
24.1362 | -35.5322 | -34.6597
24.1424 | -35.5325 | -34.6657
24.1485 | -35.5329 | -34.6716
24.1546 | -35.5333 | -34.6775
24.1607 | -35.5337 | -34.6835
24.1668 | -35.534 | -34.689
24.1729 | -35.5344 | -34.6952
24.179 | -35.5348 | -34.701
24.1851 | -35.5352 | -34.7073
24.1912 | -35.5356 | -34.7136
24.1973 | -35.536 | -34.7197
24.2034 | -35.5364 | -34.7252
24.2095 | -35.5368 | -34.7308
24.2156 | -35.5373 | -34.7367
24.2217 | -35.5377 | -34.7423
24.2278 | -35.538 | -34.7486
24.2339 | -35.5384 | -34.7546
24.24 | -35.5388 | -34.7599
24.2461 | -35.5391 | -34.7655
24.2522 | -35.5395 | -34.7713
24.2583 | -35.5399 | -34.777
24.2644 | -35.5403 | -34.783
24.2705 | -35.5407 | -34.7888
24.2766 | -35.5411 | -34.7947
24.2827 | -35.5415 | -34.8008
24.2889 | -35.5418 | -34.8068
24.295 | -35.5422 | -34.8126
24.3011 | -35.5426 | -34.8181
24.3072 | -35.543 | -34.8235
24.3133 | -35.5434 | -34.8293
24.3194 | -35.5437 | -34.8351
24.3255 | -35.544 | -34.8409
24.3316 | -35.5445 | -34.8467
24.3377 | -35.5448 | -34.8525
24.3438 | -35.5452 | -34.858
24.3499 | -35.5456 | -34.8639
24.356 | -35.546 | -34.8696
24.3621 | -35.5464 | -34.8754
24.3682 | -35.5468 | -34.8813
24.3743 | -35.5472 | -34.8873
24.3804 | -35.5475 | -34.8927
24.3865 | -35.548 | -34.8986
24.3926 | -35.5484 | -34.9048
24.3987 | -35.5487 | -34.9107
24.4048 | -35.5491 | -34.9168
24.4109 | -35.5495 | -34.9219
24.417 | -35.5499 | -34.9277
24.4231 | -35.5503 | -34.9338
24.4293 | -35.5506 | -34.9395
24.4354 | -35.551 | -34.9454
24.4415 | -35.5514 | -34.9511
24.4476 | -35.5517 | -34.9564
24.4537 | -35.5521 | -34.9624
24.4598 | -35.5525 | -34.9683
24.4659 | -35.5529 | -34.974
24.472 | -35.5533 | -34.9796
24.4781 | -35.5537 | -34.9853
24.4842 | -35.554 | -34.9909
24.4903 | -35.5543 | -34.9971
24.4964 | -35.5547 | -35.0032
24.5025 | -35.5551 | -35.009
24.5086 | -35.5554 | -35.015
24.5147 | -35.5558 | -35.0209
24.5208 | -35.5563 | -35.0267
24.5269 | -35.5566 | -35.0323
24.533 | -35.5571 | -35.0383
24.5391 | -35.5575 | -35.0442
24.5452 | -35.5578 | -35.0497
24.5513 | -35.5582 | -35.0551
24.5574 | -35.5586 | -35.0608
24.5635 | -35.5589 | -35.0668
24.5696 | -35.5593 | -35.0719
24.5758 | -35.5597 | -35.0778
24.5819 | -35.5601 | -35.0837
24.588 | -35.5605 | -35.0895
24.5941 | -35.5609 | -35.0949
24.6002 | -35.5613 | -35.1006
24.6063 | -35.5617 | -35.1062
24.6124 | -35.562 | -35.1122
24.6185 | -35.5624 | -35.1177
24.6246 | -35.5628 | -35.1227
24.6307 | -35.5631 | -35.1284
24.6368 | -35.5636 | -35.1343
24.6429 | -35.5639 | -35.1402
24.649 | -35.5643 | -35.1463
24.6551 | -35.5646 | -35.1523
24.6612 | -35.5649 | -35.1573
24.6673 | -35.5653 | -35.1631
24.6734 | -35.5657 | -35.1689
24.6795 | -35.5661 | -35.1744
24.6856 | -35.5665 | -35.18
24.6917 | -35.5669 | -35.1855
24.6978 | -35.5673 | -35.1904
24.7039 | -35.5676 | -35.1963
24.71 | -35.5679 | -35.2022
24.7162 | -35.5684 | -35.208
24.7223 | -35.5687 | -35.2138
24.7284 | -35.5691 | -35.2193
24.7345 | -35.5694 | -35.2248
24.7406 | -35.5698 | -35.2303
24.7467 | -35.5701 | -35.2357
24.7528 | -35.5706 | -35.2415
24.7589 | -35.571 | -35.247
24.765 | -35.5714 | -35.2521
24.7711 | -35.5717 | -35.2576
24.7772 | -35.5722 | -35.2632
24.7833 | -35.5726 | -35.2691
24.7894 | -35.573 | -35.2746
24.7955 | -35.5733 | -35.2805
24.8016 | -35.5737 | -35.2857
24.8077 | -35.5741 | -35.2911
24.8138 | -35.5745 | -35.2967
24.8199 | -35.5749 | -35.3023
24.826 | -35.5753 | -35.3081
24.8321 | -35.5757 | -35.3139
24.8382 | -35.576 | -35.3188
24.8443 | -35.5763 | -35.3247
24.8504 | -35.5767 | -35.3306
24.8565 | -35.5771 | -35.336
24.8627 | -35.5775 | -35.3416
24.8688 | -35.5779 | -35.3469
24.8749 | -35.5783 | -35.3527
24.881 | -35.5786 | -35.3586
24.8871 | -35.579 | -35.3642
24.8932 | -35.5794 | -35.3696
24.8993 | -35.5798 | -35.3751
24.9054 | -35.5802 | -35.3807
24.9115 | -35.5805 | -35.3865
24.9176 | -35.5809 | -35.3923
24.9237 | -35.5813 | -35.3977
24.9298 | -35.5817 | -35.4035
24.9359 | -35.582 | -35.4093
24.942 | -35.5823 | -35.4146
24.9481 | -35.5827 | -35.4203
24.9542 | -35.5831 | -35.4258
24.9603 | -35.5835 | -35.4312
24.9664 | -35.5838 | -35.4368
24.9725 | -35.5842 | -35.4428
24.9786 | -35.5845 | -35.4478
24.9847 | -35.5849 | -35.4533
24.9908 | -35.5853 | -35.4587
24.9969 | -35.5857 | -35.4644
25.0031 | -35.586 | -35.4698
25.0092 | -35.5864 | -35.4751
25.0153 | -35.5868 | -35.481
25.0214 | -35.5872 | -35.4863
25.0275 | -35.5876 | -35.4921
25.0336 | -35.588 | -35.4976
25.0397 | -35.5883 | -35.5031
25.0458 | -35.5887 | -35.5087
25.0519 | -35.5891 | -35.5143
25.058 | -35.5894 | -35.5199
25.0641 | -35.5898 | -35.5252
25.0702 | -35.5901 | -35.5307
25.0763 | -35.5905 | -35.5363
25.0824 | -35.5908 | -35.5417
25.0885 | -35.5912 | -35.5475
25.0946 | -35.5915 | -35.5531
25.1007 | -35.5918 | -35.5588
25.1068 | -35.5922 | -35.5641
25.1129 | -35.5926 | -35.5698
25.119 | -35.593 | -35.5745
25.1251 | -35.5934 | -35.58
25.1312 | -35.5937 | -35.5857
25.1373 | -35.5941 | -35.5911
25.1435 | -35.5945 | -35.5967
25.1496 | -35.5949 | -35.6022
25.1557 | -35.5953 | -35.608
25.1618 | -35.5956 | -35.6132
25.1679 | -35.596 | -35.6188
25.174 | -35.5963 | -35.6244
25.1801 | -35.5967 | -35.6299
25.1862 | -35.5971 | -35.6351
25.1923 | -35.5974 | -35.6409
25.1984 | -35.5979 | -35.6465
25.2045 | -35.5982 | -35.652
25.2106 | -35.5986 | -35.6576
25.2167 | -35.599 | -35.663
25.2228 | -35.5993 | -35.6682
25.2289 | -35.5997 | -35.6736
25.235 | -35.6001 | -35.6792
25.2411 | -35.6004 | -35.6849
25.2472 | -35.6008 | -35.6907
25.2533 | -35.6011 | -35.6964
25.2594 | -35.6015 | -35.7018
25.2655 | -35.6019 | -35.7073
25.2716 | -35.6022 | -35.713
25.2777 | -35.6025 | -35.7187
25.2838 | -35.6029 | -35.7241
25.29 | -35.6033 | -35.7296
25.2961 | -35.6036 | -35.7348
25.3022 | -35.604 | -35.7402
25.3083 | -35.6044 | -35.7456
25.3144 | -35.6048 | -35.7512
25.3205 | -35.6052 | -35.7566
25.3266 | -35.6055 | -35.7613
25.3327 | -35.6058 | -35.767
25.3388 | -35.6061 | -35.7727
25.3449 | -35.6065 | -35.7781
25.351 | -35.6069 | -35.7835
25.3571 | -35.6073 | -35.7889
25.3632 | -35.6076 | -35.794
25.3693 | -35.608 | -35.7991
25.3754 | -35.6084 | -35.8045
25.3815 | -35.6087 | -35.8101
25.3876 | -35.6091 | -35.8157
25.3937 | -35.6095 | -35.8214
25.3998 | -35.6099 | -35.8264
25.4059 | -35.6103 | -35.8318
25.412 | -35.6107 | -35.8371
25.4181 | -35.611 | -35.8427
25.4242 | -35.6114 | -35.848
25.4304 | -35.6118 | -35.8535
25.4365 | -35.6121 | -35.8587
25.4426 | -35.6125 | -35.8641
25.4487 | -35.6128 | -35.8692
25.4548 | -35.6132 | -35.8749
25.4609 | -35.6135 | -35.8801
25.467 | -35.6138 | -35.8852
25.4731 | -35.6142 | -35.8908
25.4792 | -35.6146 | -35.8963
25.4853 | -35.615 | -35.9015
25.4914 | -35.6153 | -35.9069
25.4975 | -35.6157 | -35.9124
25.5036 | -35.6159 | -35.9173
25.5097 | -35.6163 | -35.9231
25.5158 | -35.6167 | -35.9286
25.5219 | -35.617 | -35.9338
25.528 | -35.6174 | -35.9392
25.5341 | -35.6177 | -35.9446
25.5402 | -35.618 | -35.9496
25.5463 | -35.6184 | -35.9548
25.5524 | -35.6188 | -35.9601
25.5585 | -35.6191 | -35.9655
25.5646 | -35.6195 | -35.9706
25.5707 | -35.6199 | -35.9761
25.5769 | -35.6202 | -35.9814
25.583 | -35.6206 | -35.9867
25.5891 | -35.6209 | -35.9919
25.5952 | -35.6213 | -35.9972
25.6013 | -35.6217 | -36.0029
25.6074 | -35.622 | -36.0078
25.6135 | -35.6223 | -36.013
25.6196 | -35.6227 | -36.0184
25.6257 | -35.6231 | -36.0239
25.6318 | -35.6234 | -36.0292
25.6379 | -35.6238 | -36.0348
25.644 | -35.6241 | -36.04
25.6501 | -35.6245 | -36.0454
25.6562 | -35.6248 | -36.0506
25.6623 | -35.6252 | -36.056
25.6684 | -35.6255 | -36.0612
25.6745 | -35.6259 | -36.0666
25.6806 | -35.6263 | -36.0716
25.6867 | -35.6266 | -36.0772
25.6928 | -35.627 | -36.0825
25.6989 | -35.6273 | -36.0879
25.705 | -35.6277 | -36.0932
25.7111 | -35.6281 | -36.0986
25.7173 | -35.6284 | -36.1038
25.7234 | -35.6287 | -36.1089
25.7295 | -35.6291 | -36.1142
25.7356 | -35.6295 | -36.1196
25.7417 | -35.6298 | -36.1247
25.7478 | -35.6302 | -36.1297
25.7539 | -35.6305 | -36.1351
25.76 | -35.6309 | -36.1405
25.7661 | -35.6312 | -36.146
25.7722 | -35.6316 | -36.1515
25.7783 | -35.632 | -36.1566
25.7844 | -35.6324 | -36.1619
25.7905 | -35.6327 | -36.1674
25.7966 | -35.6331 | -36.173
25.8027 | -35.6334 | -36.1783
25.8088 | -35.6338 | -36.1837
25.8149 | -35.6341 | -36.1892
25.821 | -35.6344 | -36.1943
25.8271 | -35.6348 | -36.1997
25.8332 | -35.6352 | -36.2049
25.8393 | -35.6356 | -36.2106
25.8454 | -35.636 | -36.216
25.8515 | -35.6363 | -36.2214
25.8576 | -35.6367 | -36.2268
25.8638 | -35.637 | -36.2321
25.8699 | -35.6374 | -36.2372
25.876 | -35.6377 | -36.2424
25.8821 | -35.638 | -36.2479
25.8882 | -35.6384 | -36.2529
25.8943 | -35.6387 | -36.2583
25.9004 | -35.6391 | -36.2638
25.9065 | -35.6395 | -36.2695
25.9126 | -35.6398 | -36.2747
25.9187 | -35.6401 | -36.28
25.9248 | -35.6404 | -36.2849
25.9309 | -35.6408 | -36.2904
25.937 | -35.6411 | -36.2957
25.9431 | -35.6414 | -36.3011
25.9492 | -35.6418 | -36.3065
25.9553 | -35.6421 | -36.3118
25.9614 | -35.6425 | -36.3169
25.9675 | -35.6428 | -36.3222
25.9736 | -35.6432 | -36.3274
25.9797 | -35.6435 | -36.3327
25.9858 | -35.6438 | -36.3381
25.9919 | -35.6442 | -36.3434
25.998 | -35.6445 | -36.3486
26.0042 | -35.6448 | -36.3538
26.0103 | -35.6451 | -36.3591
26.0164 | -35.6455 | -36.3646
26.0225 | -35.6458 | -36.3699
26.0286 | -35.6461 | -36.3753
26.0347 | -35.6464 | -36.38
26.0408 | -35.6468 | -36.3855
26.0469 | -35.6471 | -36.3908
26.053 | -35.6475 | -36.3964
26.0591 | -35.6478 | -36.4017
26.0652 | -35.6482 | -36.4069
26.0713 | -35.6485 | -36.4121
26.0774 | -35.6488 | -36.4175
26.0835 | -35.6492 | -36.423
26.0896 | -35.6495 | -36.4284
26.0957 | -35.6499 | -36.4336
26.1018 | -35.6502 | -36.4388
26.1079 | -35.6506 | -36.4439
26.114 | -35.6509 | -36.4494
26.1201 | -35.6512 | -36.4546
26.1262 | -35.6516 | -36.46
26.1323 | -35.652 | -36.4651
26.1384 | -35.6523 | -36.47
26.1445 | -35.6527 | -36.4753
26.1507 | -35.6531 | -36.4806
26.1568 | -35.6534 | -36.4859
26.1629 | -35.6537 | -36.4911
26.169 | -35.6541 | -36.4963
26.1751 | -35.6544 | -36.5015
26.1812 | -35.6547 | -36.5068
26.1873 | -35.655 | -36.5123
26.1934 | -35.6553 | -36.5175
26.1995 | -35.6557 | -36.5226
26.2056 | -35.656 | -36.5277
26.2117 | -35.6563 | -36.5331
26.2178 | -35.6567 | -36.5385
26.2239 | -35.657 | -36.5438
26.23 | -35.6574 | -36.549
26.2361 | -35.6577 | -36.5544
26.2422 | -35.6581 | -36.5595
26.2483 | -35.6584 | -36.5645
26.2544 | -35.6587 | -36.5697
26.2605 | -35.6591 | -36.5749
26.2666 | -35.6594 | -36.5802
26.2727 | -35.6598 | -36.5855
26.2788 | -35.6601 | -36.5905
26.2849 | -35.6604 | -36.5958
26.2911 | -35.6608 | -36.601
26.2972 | -35.6611 | -36.6061
26.3033 | -35.6615 | -36.6114
26.3094 | -35.6618 | -36.6168
26.3155 | -35.6621 | -36.6214
26.3216 | -35.6624 | -36.6266
26.3277 | -35.6628 | -36.6318
26.3338 | -35.6631 | -36.6372
26.3399 | -35.6635 | -36.6423
26.346 | -35.6638 | -36.6471
26.3521 | -35.6641 | -36.6525
26.3582 | -35.6645 | -36.6577
26.3643 | -35.6648 | -36.6627
26.3704 | -35.6651 | -36.668
26.3765 | -35.6655 | -36.6733
26.3826 | -35.6658 | -36.6785
26.3887 | -35.6661 | -36.6837
26.3948 | -35.6665 | -36.689
26.4009 | -35.6668 | -36.6942
26.407 | -35.6671 | -36.6995
26.4131 | -35.6675 | -36.7047
26.4192 | -35.6678 | -36.7096
26.4253 | -35.6681 | -36.7149
26.4314 | -35.6684 | -36.7203
26.4376 | -35.6688 | -36.7253
26.4437 | -35.6691 | -36.7306
26.4498 | -35.6694 | -36.736
26.4559 | -35.6697 | -36.7409
26.462 | -35.6701 | -36.7458
26.4681 | -35.6705 | -36.751
26.4742 | -35.6708 | -36.7561
26.4803 | -35.6711 | -36.7611
26.4864 | -35.6715 | -36.7659
26.4925 | -35.6718 | -36.7709
26.4986 | -35.6721 | -36.7764
26.5047 | -35.6725 | -36.7813
26.5108 | -35.6728 | -36.7866
26.5169 | -35.6731 | -36.7919
26.523 | -35.6735 | -36.7966
26.5291 | -35.6738 | -36.8018
26.5352 | -35.6742 | -36.8069
26.5413 | -35.6744 | -36.8122
26.5474 | -35.6748 | -36.8173
26.5535 | -35.6751 | -36.8227
26.5596 | -35.6754 | -36.8278
26.5657 | -35.6758 | -36.8327
26.5718 | -35.676 | -36.8377
26.578 | -35.6764 | -36.8428
26.5841 | -35.6767 | -36.8477
26.5902 | -35.677 | -36.8529
26.5963 | -35.6773 | -36.8578
26.6024 | -35.6777 | -36.863
26.6085 | -35.678 | -36.8683
26.6146 | -35.6784 | -36.8737
26.6207 | -35.6787 | -36.8785
26.6268 | -35.679 | -36.8835
26.6329 | -35.6792 | -36.889
26.639 | -35.6796 | -36.8942
26.6451 | -35.6799 | -36.8992
26.6512 | -35.6802 | -36.9044
26.6573 | -35.6806 | -36.9094
26.6634 | -35.6808 | -36.9145
26.6695 | -35.6812 | -36.9199
26.6756 | -35.6815 | -36.9248
26.6817 | -35.6819 | -36.9299
26.6878 | -35.6822 | -36.9351
26.6939 | -35.6825 | -36.9405
26.7 | -35.6829 | -36.9455
26.7061 | -35.6831 | -36.9503
26.7122 | -35.6834 | -36.9558
26.7183 | -35.6837 | -36.9607
26.7245 | -35.6841 | -36.9661
26.7306 | -35.6844 | -36.9713
26.7367 | -35.6848 | -36.9763
26.7428 | -35.6851 | -36.9815
26.7489 | -35.6854 | -36.9866
26.755 | -35.6857 | -36.9918
26.7611 | -35.6861 | -36.9969
26.7672 | -35.6864 | -37.002
26.7733 | -35.6868 | -37.0072
26.7794 | -35.6871 | -37.0123
26.7855 | -35.6874 | -37.0175
26.7916 | -35.6877 | -37.0226
26.7977 | -35.6881 | -37.0277
26.8038 | -35.6884 | -37.0323
26.8099 | -35.6887 | -37.0377
26.816 | -35.689 | -37.0429
26.8221 | -35.6894 | -37.0481
26.8282 | -35.6897 | -37.0535
26.8343 | -35.69 | -37.0591
26.8404 | -35.6903 | -37.0641
26.8465 | -35.6906 | -37.0689
26.8526 | -35.6909 | -37.0743
26.8587 | -35.6912 | -37.0792
26.8649 | -35.6915 | -37.0843
26.871 | -35.6918 | -37.0894
26.8771 | -35.6921 | -37.0945
26.8832 | -35.6925 | -37.0997
26.8893 | -35.6928 | -37.1047
26.8954 | -35.6931 | -37.1096
26.9015 | -35.6934 | -37.1147
26.9076 | -35.6938 | -37.1198
26.9137 | -35.6941 | -37.1247
26.9198 | -35.6945 | -37.1301
26.9259 | -35.6948 | -37.1352
26.932 | -35.6951 | -37.1402
26.9381 | -35.6954 | -37.1451
26.9442 | -35.6958 | -37.1503
26.9503 | -35.6961 | -37.1553
26.9564 | -35.6964 | -37.1606
26.9625 | -35.6967 | -37.1657
26.9686 | -35.6971 | -37.1707
26.9747 | -35.6974 | -37.1756
26.9808 | -35.6978 | -37.1805
26.9869 | -35.698 | -37.1857
26.993 | -35.6984 | -37.1909
26.9991 | -35.6987 | -37.1961
27.0052 | -35.699 | -37.2013
27.0114 | -35.6994 | -37.2066
27.0175 | -35.6997 | -37.2115
27.0236 | -35.7 | -37.2165
27.0297 | -35.7003 | -37.2215
27.0358 | -35.7007 | -37.2264
27.0419 | -35.701 | -37.2313
27.048 | -35.7014 | -37.2365
27.0541 | -35.7017 | -37.2418
27.0602 | -35.702 | -37.2466
27.0663 | -35.7023 | -37.2522
27.0724 | -35.7027 | -37.2572
27.0785 | -35.703 | -37.2627
27.0846 | -35.7033 | -37.2676
27.0907 | -35.7036 | -37.2725
27.0968 | -35.7039 | -37.2777
27.1029 | -35.7042 | -37.2828
27.109 | -35.7045 | -37.288
27.1151 | -35.7048 | -37.2934
27.1212 | -35.7052 | -37.2986
27.1273 | -35.7055 | -37.3034
27.1334 | -35.7059 | -37.3086
27.1395 | -35.7062 | -37.314
27.1456 | -35.7064 | -37.3189
27.1518 | -35.7068 | -37.3237
27.1579 | -35.7071 | -37.3289
27.164 | -35.7074 | -37.3342
27.1701 | -35.7077 | -37.3392
27.1762 | -35.708 | -37.3443
27.1823 | -35.7083 | -37.3495
27.1884 | -35.7086 | -37.3547
27.1945 | -35.709 | -37.3598
27.2006 | -35.7093 | -37.365
27.2067 | -35.7096 | -37.3704
27.2128 | -35.71 | -37.3755
27.2189 | -35.7103 | -37.3807
27.225 | -35.7106 | -37.3856
27.2311 | -35.7109 | -37.3905
27.2372 | -35.7112 | -37.3956
27.2433 | -35.7115 | -37.4007
27.2494 | -35.7118 | -37.4055
27.2555 | -35.7122 | -37.4108
27.2616 | -35.7125 | -37.4157
27.2677 | -35.7128 | -37.4204
27.2738 | -35.7131 | -37.4258
27.2799 | -35.7135 | -37.4306
27.286 | -35.7138 | -37.4356
27.2921 | -35.7141 | -37.4408
27.2983 | -35.7144 | -37.4457
27.3044 | -35.7148 | -37.4509
27.3105 | -35.7151 | -37.4561
27.3166 | -35.7154 | -37.4611
27.3227 | -35.7157 | -37.4662
27.3288 | -35.716 | -37.4712
27.3349 | -35.7163 | -37.4766
27.341 | -35.7167 | -37.4822
27.3471 | -35.717 | -37.4872
27.3532 | -35.7173 | -37.4919
27.3593 | -35.7176 | -37.4968
27.3654 | -35.718 | -37.5014
27.3715 | -35.7183 | -37.5063
27.3776 | -35.7186 | -37.5114
27.3837 | -35.7189 | -37.5167
27.3898 | -35.7192 | -37.5221
27.3959 | -35.7195 | -37.5274
27.402 | -35.7198 | -37.532
27.4081 | -35.7201 | -37.5374
27.4142 | -35.7204 | -37.5424
27.4203 | -35.7207 | -37.5477
27.4264 | -35.721 | -37.5529
27.4325 | -35.7214 | -37.5581
27.4387 | -35.7217 | -37.5627
27.4448 | -35.722 | -37.568
27.4509 | -35.7223 | -37.573
27.457 | -35.7226 | -37.5779
27.4631 | -35.7229 | -37.5832
27.4692 | -35.7233 | -37.5882
27.4753 | -35.7236 | -37.5932
27.4814 | -35.7239 | -37.5984
27.4875 | -35.7242 | -37.6036
27.4936 | -35.7245 | -37.6088
27.4997 | -35.7249 | -37.6138
27.5058 | -35.7252 | -37.6187
27.5119 | -35.7255 | -37.6239
27.518 | -35.7259 | -37.6286
27.5241 | -35.7262 | -37.6334
27.5302 | -35.7265 | -37.6385
27.5363 | -35.7267 | -37.6435
27.5424 | -35.727 | -37.6486
27.5485 | -35.7273 | -37.6536
27.5546 | -35.7277 | -37.6588
27.5607 | -35.728 | -37.6635
27.5668 | -35.7283 | -37.6685
27.5729 | -35.7286 | -37.6736
27.5791 | -35.7289 | -37.6784
27.5852 | -35.7292 | -37.6836
27.5913 | -35.7295 | -37.6884
27.5974 | -35.7298 | -37.6933
27.6035 | -35.7302 | -37.6981
27.6096 | -35.7305 | -37.7031
27.6157 | -35.7308 | -37.7083
27.6218 | -35.7311 | -37.7134
27.6279 | -35.7314 | -37.7184
27.634 | -35.7318 | -37.7234
27.6401 | -35.732 | -37.7285
27.6462 | -35.7324 | -37.7334
27.6523 | -35.7327 | -37.7388
27.6584 | -35.733 | -37.7437
27.6645 | -35.7333 | -37.7488
27.6706 | -35.7336 | -37.754
27.6767 | -35.7339 | -37.759
27.6828 | -35.7342 | -37.7639
27.6889 | -35.7345 | -37.7691
27.695 | -35.7348 | -37.7739
27.7011 | -35.7351 | -37.7789
27.7072 | -35.7354 | -37.784
27.7133 | -35.7356 | -37.7889
27.7194 | -35.736 | -37.7939
27.7256 | -35.7364 | -37.7987
27.7317 | -35.7366 | -37.8036
27.7378 | -35.737 | -37.8086
27.7439 | -35.7372 | -37.8136
27.75 | -35.7376 | -37.8182
27.7561 | -35.7378 | -37.8234
27.7622 | -35.7381 | -37.8282
27.7683 | -35.7385 | -37.8332
27.7744 | -35.7388 | -37.838
27.7805 | -35.739 | -37.8431
27.7866 | -35.7393 | -37.8485
27.7927 | -35.7396 | -37.853
27.7988 | -35.7399 | -37.858
27.8049 | -35.7402 | -37.8627
27.811 | -35.7406 | -37.8675
27.8171 | -35.7409 | -37.8728
27.8232 | -35.7412 | -37.8779
27.8293 | -35.7415 | -37.8829
27.8354 | -35.7418 | -37.8879
27.8415 | -35.7422 | -37.8932
27.8476 | -35.7424 | -37.8983
27.8537 | -35.7427 | -37.9031
27.8598 | -35.743 | -37.9081
27.866 | -35.7434 | -37.9134
27.8721 | -35.7437 | -37.9184
27.8782 | -35.744 | -37.9237
27.8843 | -35.7442 | -37.9289
27.8904 | -35.7445 | -37.9334
27.8965 | -35.7447 | -37.9383
27.9026 | -35.7451 | -37.9431
27.9087 | -35.7454 | -37.948
27.9148 | -35.7457 | -37.9534
27.9209 | -35.746 | -37.9585
27.927 | -35.7463 | -37.9638
27.9331 | -35.7466 | -37.9693
27.9392 | -35.7468 | -37.9741
27.9453 | -35.7471 | -37.979
27.9514 | -35.7474 | -37.9845
27.9575 | -35.7477 | -37.9896
27.9636 | -35.748 | -37.9949
27.9697 | -35.7484 | -37.9998
27.9758 | -35.7488 | -38.0045
27.9819 | -35.749 | -38.0092
27.988 | -35.7494 | -38.0144
27.9941 | -35.7497 | -38.0194
28.0002 | -35.7499 | -38.0244
28.0063 | -35.7501 | -38.0291
28.0125 | -35.7504 | -38.0342
28.0186 | -35.7507 | -38.0392
28.0247 | -35.7509 | -38.0443
28.0308 | -35.7512 | -38.0492
28.0369 | -35.7514 | -38.0539
28.043 | -35.7517 | -38.0591
28.0491 | -35.7519 | -38.064
28.0552 | -35.7523 | -38.0692
28.0613 | -35.7526 | -38.0738
28.0674 | -35.7528 | -38.0779
28.0735 | -35.753 | -38.0831

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.477 | -42.1953
21.4908 | -35.5978 | -35.1922
16.4948 | -34.9818 | -27.7982
12.6602 | -34.6249 | -21.5218
9.7171 | -34.4059 | -16.608
7.4582 | -34.2528 | -12.6221
5.7244 | -34.152 | -9.3928
4.3936 | -34.0825 | -6.7852
3.3722 | -34.0628 | -4.7332
2.5883 | -34.0593 | -3.1728
1.9866 | -34.0481 | -1.9285
1.5248 | -34.0512 | -0.92918
1.1703 | -34.0572 | -0.27903
1 | -34.0553 | 0.042314
0.89822 | -34.053 | 0.31482
0.68941 | -34.0346 | 0.89554
0.52919 | -34.0393 | 1.347
0.40612 | -34.0376 | 1.9535
0.31171 | -34.0282 | 2.4052
0.23928 | -34.0275 | 3.132
0.18358 | -34.0405 | 3.5499
0.14091 | -34.063 | 4.4239
0.10819 | -34.0581 | 5.5926
0.083017 | -34.0837 | 7.0676
0.063705 | -34.1476 | 8.9233
0.048923 | -34.2121 | 11.3986
0.037479 | -34.3082 | 14.584
0.028801 | -34.4788 | 18.6098
0.022078 | -34.7903 | 24.0514
0.017023 | -35.2824 | 30.5003
0.013018 | -36.0833 | 38.1626
0.010014 | -37.2053 | 46.07

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003719 | 0.37679
0.007438 | -27.9294
0.011158 | -35.1472
0.014877 | -41.7121
0.018596 | -43.1409
0.022315 | -44.3319
0.026034 | -46.4706
0.029754 | -47.5513
0.033473 | -49.9871
0.037192 | -51.1669
0.040911 | -52.8301
0.04463 | -53.1699
0.04835 | -53.669
0.052069 | -54.3842
0.055788 | -55.4438
0.059507 | -56.2031
0.063226 | -57.307
0.066946 | -57.7815
0.070665 | -57.4645
0.074384 | -58.3943
0.078103 | -58.7483
0.081822 | -59.9445
0.085542 | -60.1901
0.089261 | -60.4351
0.09298 | -61.2742
0.096699 | -61.3623
0.10042 | -61.9826
0.10414 | -62.5013
0.10786 | -62.1854
0.11158 | -62.1007
0.11529 | -62.6443
0.11901 | -63.9525
0.12273 | -64.119
0.12645 | -65.9095
0.13017 | -67.0733
0.13389 | -65.5926
0.13761 | -65.9758
0.14133 | -66.4347
0.14505 | -66.6156
0.14877 | -66.0919
0.15249 | -65.7946
0.15621 | -66.7901
0.15993 | -67.2294
0.16365 | -68.074
0.16736 | -68.0569
0.17108 | -67.2486
0.1748 | -67.7566
0.17852 | -67.4993
0.18224 | -68.1432
0.18596 | -67.2933
0.18968 | -67.97
0.1934 | -68.2099
0.19712 | -67.5721
0.20084 | -67.8911
0.20456 | -69.0483
0.20827 | -69.6185
0.21199 | -69.7512
0.21571 | -69.9687
0.21943 | -70.4564
0.22315 | -71.526
0.22687 | -70.5751
0.23059 | -69.726
0.23431 | -69.2059
0.23803 | -69.5823
0.24175 | -69.9196
0.24547 | -71.2205
0.24919 | -71.0653
0.25291 | -70.9759
0.25662 | -70.7006
0.26034 | -72.0402
0.26406 | -71.6367
0.26778 | -71.6391
0.2715 | -72.5583
0.27522 | -73.7509
0.27894 | -73.6293
0.28266 | -72.5126
0.28638 | -72.1783
0.2901 | -72.9665
0.29382 | -73.0428
0.29754 | -72.9789
0.30125 | -73.6802
0.30497 | -73.4353
0.30869 | -74.2855
0.31241 | -73.6952
0.31613 | -74.0649
0.31985 | -73.9931
0.32357 | -74.7194
0.32729 | -74.8152
0.33101 | -74.2407
0.33473 | -73.7984
0.33845 | -74.1887
0.34217 | -74.291
0.34589 | -75.5532
0.3496 | -75.4996
0.35332 | -74.9264
0.35704 | -75.9891
0.36076 | -76.3282
0.36448 | -76.5387
0.3682 | -76.7008
0.37192 | -76.1827
0.37564 | -75.0761
0.37936 | -75.3262
0.38308 | -75.4696
0.3868 | -75.9699
0.39052 | -76.0126
0.39424 | -76.7455
0.39795 | -78.6248
0.40167 | -77.5439
0.40539 | -77.7462
0.40911 | -77.6107
0.41283 | -77.4294
0.41655 | -77.5793
0.42027 | -77.6253
0.42399 | -77.9544
0.42771 | -77.6581
0.43143 | -78.0856
0.43515 | -77.7515
0.43887 | -77.7057
0.44259 | -78.325
0.4463 | -77.3678
0.45002 | -77.8303
0.45374 | -78.6541
0.45746 | -78.0188
0.46118 | -77.7241
0.4649 | -78.7319
0.46862 | -77.9261
0.47234 | -77.7686
0.47606 | -77.9084
0.47978 | -78.5198
0.4835 | -79.0827
0.48722 | -79.6916
0.49093 | -79.7796
0.49465 | -78.8375
0.49837 | -78.6644
0.50209 | -79.6572
0.50581 | -80.4025
0.50953 | -79.7238
0.51325 | -80.2881
0.51697 | -79.9876
0.52069 | -80.2434
0.52441 | -79.7967
0.52813 | -79.9205
0.53185 | -80.0118
0.53556 | -79.1821
0.53928 | -79.0638
0.543 | -79.567
0.54672 | -80.2687
0.55044 | -80.5661
0.55416 | -80.6309
0.55788 | -80.4835
0.5616 | -79.9595
0.56532 | -79.6195
0.56904 | -79.8537
0.57276 | -80.0057
0.57648 | -79.8717
0.5802 | -80.1405
0.58391 | -80.6168
0.58763 | -80.7188
0.59135 | -80.5407
0.59507 | -80.507
0.59879 | -81.1061
0.60251 | -81.6642
0.60623 | -81.8271
0.60995 | -82.2191
0.61367 | -82.4862
0.61739 | -81.9049
0.62111 | -81.3485
0.62483 | -81.4548
0.62855 | -81.759
0.63226 | -81.9997
0.63598 | -82.0701
0.6397 | -81.7361
0.64342 | -81.2716
0.64714 | -81.2021
0.65086 | -81.2918
0.65458 | -81.6372
0.6583 | -82.4159
0.66202 | -82.6249
0.66574 | -82.1835
0.66946 | -81.9435
0.67317 | -82.307
0.67689 | -82.6502
0.68061 | -82.6274
0.68433 | -82.6153
0.68805 | -82.8278
0.69177 | -82.7188
0.69549 | -82.5122
0.69921 | -82.548
0.70293 | -83.0879
0.70665 | -83.8848
0.71037 | -83.5179
0.71409 | -82.6159
0.71781 | -82.5445
0.72152 | -82.9144
0.72524 | -82.9222
0.72896 | -82.7838
0.73268 | -82.8122
0.7364 | -82.6801
0.74012 | -82.6215
0.74384 | -82.9211
0.74756 | -83.1198
0.75128 | -82.9097
0.755 | -82.9885
0.75872 | -83.5983
0.76244 | -84.206
0.76616 | -84.5888
0.76987 | -84.6787
0.77359 | -84.2442
0.77731 | -83.4205
0.78103 | -82.9727
0.78475 | -83.0042
0.78847 | -83.3551
0.79219 | -84.2427
0.79591 | -85.0631
0.79963 | -84.9434
0.80335 | -84.1832
0.80707 | -83.7537
0.81079 | -83.8522
0.81451 | -84.2305
0.81822 | -84.4873
0.82194 | -84.4521
0.82566 | -84.1339
0.82938 | -83.9029
0.8331 | -83.6061
0.83682 | -83.4727
0.84054 | -83.5868
0.84426 | -83.7284
0.84798 | -84.1603
0.8517 | -84.5721
0.85542 | -84.9075
0.85913 | -85.4691
0.86285 | -85.8848
0.86657 | -85.8327
0.87029 | -85.8543
0.87401 | -86.0164
0.87773 | -85.5995
0.88145 | -85.0553
0.88517 | -85.0578
0.88889 | -85.1876
0.89261 | -85.3297
0.89633 | -85.6971
0.90005 | -86.1413
0.90377 | -86.2964
0.90748 | -86.0571
0.9112 | -85.4564
0.91492 | -85.4784
0.91864 | -86.0811
0.92236 | -86.0768
0.92608 | -85.5517
0.9298 | -84.7588
0.93352 | -84.5206
0.93724 | -84.8791
0.94096 | -85.3639
0.94468 | -85.9268
0.9484 | -86.7469
0.95212 | -87.1166
0.95583 | -86.9931
0.95955 | -86.8883
0.96327 | -86.9069
0.96699 | -86.5964
0.97071 | -86.2932
0.97443 | -86.2329
0.97815 | -86.3163
0.98187 | -86.4107
0.98559 | -86.809
0.98931 | -87.3076
0.99303 | -86.9025
0.99675 | -85.9859
1.0005 | -85.5575
1.0042 | -85.8859
1.0079 | -86.3255
1.0116 | -86.7969
1.0153 | -87.4662
1.0191 | -87.9631
1.0228 | -87.5934
1.0265 | -86.9863
1.0302 | -86.8475
1.0339 | -86.9467
1.0377 | -86.6776
1.0414 | -86.6364
1.0451 | -86.9979
1.0488 | -87.5253
1.0525 | -87.9154
1.0563 | -87.6856
1.06 | -87.1355
1.0637 | -86.8652
1.0674 | -86.8889
1.0711 | -86.9341
1.0748 | -87.0035
1.0786 | -87.0554
1.0823 | -87.0694
1.086 | -87.1192
1.0897 | -87.2337
1.0934 | -87.3808
1.0972 | -87.4955
1.1009 | -87.4954
1.1046 | -87.3158
1.1083 | -87.089
1.112 | -87.1226
1.1158 | -87.5112
1.1195 | -88.0087
1.1232 | -88.3718
1.1269 | -88.5161
1.1306 | -88.3556
1.1344 | -87.9328
1.1381 | -87.4567
1.1418 | -87.1084
1.1455 | -87.0444
1.1492 | -87.2583
1.153 | -87.4584
1.1567 | -87.443
1.1604 | -87.4401
1.1641 | -87.733
1.1678 | -88.2305
1.1715 | -88.5236
1.1753 | -88.4598
1.179 | -88.358
1.1827 | -88.4344
1.1864 | -88.5238
1.1901 | -88.413
1.1939 | -88.1969
1.1976 | -88.0746
1.2013 | -88.0442
1.205 | -87.9964
1.2087 | -88.0046
1.2125 | -88.2159
1.2162 | -88.5147
1.2199 | -88.6403
1.2236 | -88.5477
1.2273 | -88.4412
1.2311 | -88.4714
1.2348 | -88.5687
1.2385 | -88.6667
1.2422 | -88.7642
1.2459 | -88.7041
1.2497 | -88.3435
1.2534 | -87.8736
1.2571 | -87.6099
1.2608 | -87.7469
1.2645 | -88.3541
1.2682 | -89.2674
1.272 | -89.917
1.2757 | -89.8562
1.2794 | -89.3359
1.2831 | -88.8698
1.2868 | -88.7655
1.2906 | -88.9268
1.2943 | -89.0226
1.298 | -88.9947
1.3017 | -89.0409
1.3054 | -89.2722
1.3092 | -89.5028
1.3129 | -89.4495
1.3166 | -89.1924
1.3203 | -89.038
1.324 | -89.1475
1.3278 | -89.4465
1.3315 | -89.6897
1.3352 | -89.7483
1.3389 | -89.651
1.3426 | -89.4573
1.3464 | -89.3107
1.3501 | -89.3213
1.3538 | -89.3614
1.3575 | -89.3077
1.3612 | -89.2113
1.3649 | -89.1496
1.3687 | -89.1891
1.3724 | -89.2757
1.3761 | -89.2534
1.3798 | -89.2349
1.3835 | -89.4843
1.3873 | -90.0656
1.391 | -90.7782
1.3947 | -91.2624
1.3984 | -91.302
1.4021 | -91.0704
1.4059 | -90.7936
1.4096 | -90.458
1.4133 | -90.0742
1.417 | -89.8321
1.4207 | -89.7823
1.4245 | -89.7709
1.4282 | -89.8012
1.4319 | -90.0172
1.4356 | -90.3575
1.4393 | -90.6321
1.443 | -90.8494
1.4468 | -91.0259
1.4505 | -90.996
1.4542 | -90.7111
1.4579 | -90.2871
1.4616 | -89.8872
1.4654 | -89.7432
1.4691 | -89.8981
1.4728 | -90.1168
1.4765 | -90.2304
1.4802 | -90.1587
1.484 | -89.8777
1.4877 | -89.7025
1.4914 | -89.9112
1.4951 | -90.3667
1.4988 | -90.6927
1.5026 | -90.7135
1.5063 | -90.6034
1.51 | -90.5048
1.5137 | -90.4396
1.5174 | -90.442
1.5212 | -90.4876
1.5249 | -90.5397
1.5286 | -90.6491
1.5323 | -90.9087
1.536 | -91.2614
1.5397 | -91.5006
1.5435 | -91.6575
1.5472 | -91.8635
1.5509 | -92.0197
1.5546 | -91.9469
1.5583 | -91.6457
1.5621 | -91.3604
1.5658 | -91.2319
1.5695 | -91.2324
1.5732 | -91.319
1.5769 | -91.2438
1.5807 | -90.8971
1.5844 | -90.6576
1.5881 | -90.7903
1.5918 | -91.1169
1.5955 | -91.2622
1.5993 | -91.224
1.603 | -91.3073
1.6067 | -91.4639
1.6104 | -91.5205
1.6141 | -91.4952
1.6179 | -91.4206
1.6216 | -91.3163
1.6253 | -91.2207
1.629 | -91.1944
1.6327 | -91.2554
1.6364 | -91.3241
1.6402 | -91.3358
1.6439 | -91.3053
1.6476 | -91.2985
1.6513 | -91.3358
1.655 | -91.3999
1.6588 | -91.4712
1.6625 | -91.5503
1.6662 | -91.5893
1.6699 | -91.5427
1.6736 | -91.4508
1.6774 | -91.4307
1.6811 | -91.4916
1.6848 | -91.5423
1.6885 | -91.5096
1.6922 | -91.4232
1.696 | -91.3789
1.6997 | -91.4505
1.7034 | -91.6266
1.7071 | -91.8347
1.7108 | -92.0038
1.7146 | -92.1041
1.7183 | -92.1335
1.722 | -92.1728
1.7257 | -92.2844
1.7294 | -92.4244
1.7331 | -92.5003
1.7369 | -92.4618
1.7406 | -92.3284
1.7443 | -92.1241
1.748 | -91.9016
1.7517 | -91.7467
1.7555 | -91.6764
1.7592 | -91.6263
1.7629 | -91.5585
1.7666 | -91.5637
1.7703 | -91.7596
1.7741 | -92.1181
1.7778 | -92.4739
1.7815 | -92.6529
1.7852 | -92.6422
1.7889 | -92.5599
1.7927 | -92.467
1.7964 | -92.351
1.8001 | -92.2173
1.8038 | -92.0776
1.8075 | -91.966
1.8113 | -91.9472
1.815 | -92.0516
1.8187 | -92.2221
1.8224 | -92.3924
1.8261 | -92.5405
1.8298 | -92.6045
1.8336 | -92.557
1.8373 | -92.497
1.841 | -92.5416
1.8447 | -92.689
1.8484 | -92.8894
1.8522 | -93.119
1.8559 | -93.3502
1.8596 | -93.4646
1.8633 | -93.3526
1.867 | -93.0435
1.8708 | -92.7201
1.8745 | -92.5379
1.8782 | -92.5228
1.8819 | -92.5454
1.8856 | -92.5052
1.8894 | -92.4724
1.8931 | -92.5647
1.8968 | -92.7298
1.9005 | -92.7964
1.9042 | -92.6807
1.908 | -92.4761
1.9117 | -92.3271
1.9154 | -92.3366
1.9191 | -92.5217
1.9228 | -92.8366
1.9265 | -93.1888
1.9303 | -93.4848
1.934 | -93.6719
1.9377 | -93.7127
1.9414 | -93.6324
1.9451 | -93.5222
1.9489 | -93.497
1.9526 | -93.5628
1.9563 | -93.5794
1.96 | -93.4351
1.9637 | -93.2081
1.9675 | -93.059
1.9712 | -93.0842
1.9749 | -93.2551
1.9786 | -93.4426
1.9823 | -93.4759
1.9861 | -93.345
1.9898 | -93.1995
1.9935 | -93.1742
1.9972 | -93.2455
2.0009 | -93.2945
2.0046 | -93.2431
2.0084 | -93.1397
2.0121 | -93.0693
2.0158 | -93.0613
2.0195 | -93.0895
2.0232 | -93.1298
2.027 | -93.1413
2.0307 | -93.1283
2.0344 | -93.1562
2.0381 | -93.2761
2.0418 | -93.4812
2.0456 | -93.6708
2.0493 | -93.7649
2.053 | -93.7928
2.0567 | -93.8102
2.0604 | -93.8266
2.0642 | -93.8398
2.0679 | -93.9011
2.0716 | -94.0392
2.0753 | -94.1454
2.079 | -94.1095
2.0828 | -93.929
2.0865 | -93.6761
2.0902 | -93.4224
2.0939 | -93.2144
2.0976 | -93.1324
2.1013 | -93.2287
2.1051 | -93.4947
2.1088 | -93.8622
2.1125 | -94.2282
2.1162 | -94.4946
2.1199 | -94.6279
2.1237 | -94.6464
2.1274 | -94.643
2.1311 | -94.6768
2.1348 | -94.7358
2.1385 | -94.7079
2.1423 | -94.5228
2.146 | -94.3044
2.1497 | -94.1681
2.1534 | -94.0979
2.1571 | -94.0729
2.1609 | -94.0605
2.1646 | -94.0276
2.1683 | -93.9744
2.172 | -93.9332
2.1757 | -93.9612
2.1795 | -94.0974
2.1832 | -94.3137
2.1869 | -94.5566
2.1906 | -94.7619
2.1943 | -94.9036
2.198 | -94.9868
2.2018 | -95.0264
2.2055 | -95.0372
2.2092 | -95.0168
2.2129 | -94.9438
2.2166 | -94.8157
2.2204 | -94.6447
2.2241 | -94.4597
2.2278 | -94.2663
2.2315 | -94.0614
2.2352 | -93.863
2.239 | -93.7246
2.2427 | -93.6887
2.2464 | -93.7736
2.2501 | -93.9568
2.2538 | -94.1749
2.2576 | -94.3545
2.2613 | -94.4729
2.265 | -94.5493
2.2687 | -94.6219
2.2724 | -94.7203
2.2762 | -94.832
2.2799 | -94.9095
2.2836 | -94.8974
2.2873 | -94.7945
2.291 | -94.6574
2.2947 | -94.5795
2.2985 | -94.6108
2.3022 | -94.7372
2.3059 | -94.8963
2.3096 | -95.0194
2.3133 | -95.0758
2.3171 | -95.0779
2.3208 | -95.0312
2.3245 | -94.9775
2.3282 | -94.9143
2.3319 | -94.8496
2.3357 | -94.8065
2.3394 | -94.7914
2.3431 | -94.7954
2.3468 | -94.7836
2.3505 | -94.7083
2.3543 | -94.5722
2.358 | -94.4272
2.3617 | -94.3331
2.3654 | -94.2912
2.3691 | -94.2982
2.3729 | -94.3238
2.3766 | -94.3529
2.3803 | -94.4085
2.384 | -94.5016
2.3877 | -94.6207
2.3914 | -94.7409
2.3952 | -94.844
2.3989 | -94.9186
2.4026 | -94.9496
2.4063 | -94.9462
2.41 | -94.9298
2.4138 | -94.9235
2.4175 | -94.9579
2.4212 | -95.0311
2.4249 | -95.1212
2.4286 | -95.2037
2.4324 | -95.2415
2.4361 | -95.2221
2.4398 | -95.1748
2.4435 | -95.0966
2.4472 | -94.9675
2.451 | -94.8092
2.4547 | -94.6501
2.4584 | -94.5189
2.4621 | -94.462
2.4658 | -94.4788
2.4695 | -94.5317
2.4733 | -94.5726
2.477 | -94.5583
2.4807 | -94.4921
2.4844 | -94.4412
2.4881 | -94.4591
2.4919 | -94.5605
2.4956 | -94.729
2.4993 | -94.8973
2.503 | -95.0195
2.5067 | -95.0855
2.5105 | -95.09
2.5142 | -95.0556
2.5179 | -95.0301
2.5216 | -95.0646
2.5253 | -95.1893
2.5291 | -95.3512
2.5328 | -95.4685
2.5365 | -95.4893
2.5402 | -95.423
2.5439 | -95.306
2.5477 | -95.2044
2.5514 | -95.1643
2.5551 | -95.1957
2.5588 | -95.2882
2.5625 | -95.4142
2.5662 | -95.5364
2.57 | -95.6516
2.5737 | -95.7204
2.5774 | -95.7202
2.5811 | -95.647
2.5848 | -95.521
2.5886 | -95.3747
2.5923 | -95.2328
2.596 | -95.1166
2.5997 | -95.062
2.6034 | -95.1084
2.6072 | -95.2736
2.6109 | -95.5371
2.6146 | -95.8251
2.6183 | -96.0348
2.622 | -96.1453
2.6258 | -96.1879
2.6295 | -96.1827
2.6332 | -96.125
2.6369 | -95.9845
2.6406 | -95.7657
2.6444 | -95.5314
2.6481 | -95.3492
2.6518 | -95.2625
2.6555 | -95.2852
2.6592 | -95.3829
2.6629 | -95.4924
2.6667 | -95.5798
2.6704 | -95.6365
2.6741 | -95.6524
2.6778 | -95.6396
2.6815 | -95.6966
2.6853 | -95.7686
2.689 | -95.869
2.6927 | -95.9973
2.6964 | -96.1126
2.7001 | -96.185
2.7039 | -96.2045
2.7076 | -96.1556
2.7113 | -96.0528
2.715 | -95.9313
2.7187 | -95.8244
2.7225 | -95.7505
2.7262 | -95.6957
2.7299 | -95.6537
2.7336 | -95.6305
2.7373 | -95.6362
2.7411 | -95.68
2.7448 | -95.7863
2.7485 | -95.9548
2.7522 | -96.1635
2.7559 | -96.3863
2.7596 | -96.5631
2.7634 | -96.6609
2.7671 | -96.6621
2.7708 | -96.5611
2.7745 | -96.4001
2.7782 | -96.1984
2.782 | -95.9785
2.7857 | -95.749
2.7894 | -95.5424
2.7931 | -95.4134
2.7968 | -95.3843
2.8006 | -95.4461
2.8043 | -95.5734
2.808 | -95.7321
2.8117 | -95.889
2.8154 | -96.0402
2.8192 | -96.1722
2.8229 | -96.2786
2.8266 | -96.3715
2.8303 | -96.4473
2.834 | -96.5192
2.8378 | -96.5874
2.8415 | -96.6368
2.8452 | -96.6791
2.8489 | -96.71
2.8526 | -96.7157
2.8563 | -96.6992
2.8601 | -96.6504
2.8638 | -96.5611
2.8675 | -96.4406
2.8712 | -96.3124
2.8749 | -96.1841
2.8787 | -96.0704
2.8824 | -95.9775
2.8861 | -95.9069
2.8898 | -95.8652
2.8935 | -95.8549
2.8973 | -95.8712
2.901 | -95.9111
2.9047 | -95.989
2.9084 | -96.1377
2.9121 | -96.3711
2.9159 | -96.6738
2.9196 | -96.994
2.9233 | -97.2888
2.927 | -97.4984
2.9307 | -97.584
2.9344 | -97.5357
2.9382 | -97.3834
2.9419 | -97.1757
2.9456 | -96.9779
2.9493 | -96.8316
2.953 | -96.7409
2.9568 | -96.692
2.9605 | -96.6599
2.9642 | -96.6183
2.9679 | -96.5599
2.9716 | -96.5117
2.9754 | -96.5164
2.9791 | -96.5899
2.9828 | -96.7163
2.9865 | -96.856
2.9902 | -96.9596
2.994 | -96.9971
2.9977 | -96.9737
3.0014 | -96.9263
3.0051 | -96.885
3.0088 | -96.8796
3.0126 | -96.9148
3.0163 | -96.9786
3.02 | -97.0385
3.0237 | -97.0788
3.0274 | -97.1095
3.0311 | -97.1233
3.0349 | -97.1059
3.0386 | -97.0677
3.0423 | -97.0335
3.046 | -97.0301
3.0497 | -97.0684
3.0535 | -97.1393
3.0572 | -97.2153
3.0609 | -97.2618
3.0646 | -97.2452
3.0683 | -97.1712
3.0721 | -97.0488
3.0758 | -96.8955
3.0795 | -96.7543
3.0832 | -96.6402
3.0869 | -96.5687
3.0907 | -96.5391
3.0944 | -96.536
3.0981 | -96.5541
3.1018 | -96.6101
3.1055 | -96.7162
3.1093 | -96.8793
3.113 | -97.0828
3.1167 | -97.2765
3.1204 | -97.4313
3.1241 | -97.5147
3.1278 | -97.4906
3.1316 | -97.3756
3.1353 | -97.2105
3.139 | -97.0432
3.1427 | -96.9179
3.1464 | -96.8384
3.1502 | -96.7916
3.1539 | -96.7754
3.1576 | -96.7867
3.1613 | -96.8071
3.165 | -96.8202
3.1688 | -96.8081
3.1725 | -96.7938
3.1762 | -96.8107
3.1799 | -96.8668
3.1836 | -96.9478
3.1874 | -97.0413
3.1911 | -97.1203
3.1948 | -97.178
3.1985 | -97.216
3.2022 | -97.2373
3.206 | -97.2398
3.2097 | -97.2341
3.2134 | -97.2313
3.2171 | -97.2493
3.2208 | -97.2714
3.2245 | -97.3069
3.2283 | -97.3532
3.232 | -97.4137
3.2357 | -97.4858
3.2394 | -97.5487
3.2431 | -97.588
3.2469 | -97.5938
3.2506 | -97.5779
3.2543 | -97.5591
3.258 | -97.5557
3.2617 | -97.5772
3.2655 | -97.6186
3.2692 | -97.6687
3.2729 | -97.7116
3.2766 | -97.7219
3.2803 | -97.6994
3.2841 | -97.6628
3.2878 | -97.6266
3.2915 | -97.6113
3.2952 | -97.6261
3.2989 | -97.6685
3.3026 | -97.7352
3.3064 | -97.8206
3.3101 | -97.923
3.3138 | -98.0393
3.3175 | -98.1576
3.3212 | -98.2519
3.325 | -98.3088
3.3287 | -98.3187
3.3324 | -98.2732
3.3361 | -98.1676
3.3398 | -98.0248
3.3436 | -97.8769
3.3473 | -97.7428
3.351 | -97.6494
3.3547 | -97.6166
3.3584 | -97.6423
3.3622 | -97.715
3.3659 | -97.8176
3.3696 | -97.9265
3.3733 | -98.0102
3.377 | -98.0558
3.3808 | -98.0505
3.3845 | -98.0006
3.3882 | -97.9217
3.3919 | -97.8292
3.3956 | -97.7282
3.3993 | -97.6376
3.4031 | -97.5761
3.4068 | -97.5409
3.4105 | -97.5328
3.4142 | -97.5547
3.4179 | -97.6035
3.4217 | -97.6665
3.4254 | -97.7355
3.4291 | -97.8158
3.4328 | -97.8979
3.4365 | -97.9636
3.4403 | -97.9992
3.444 | -98.0097
3.4477 | -98.0027
3.4514 | -98.0116
3.4551 | -98.0548
3.4589 | -98.1373
3.4626 | -98.2563
3.4663 | -98.3932
3.47 | -98.5247
3.4737 | -98.6276
3.4775 | -98.6903
3.4812 | -98.6898
3.4849 | -98.6231
3.4886 | -98.5148
3.4923 | -98.3994
3.496 | -98.3095
3.4998 | -98.2754
3.5035 | -98.3134
3.5072 | -98.4178
3.5109 | -98.5725
3.5146 | -98.7437
3.5184 | -98.8979
3.5221 | -98.9937
3.5258 | -99.0097
3.5295 | -98.9417
3.5332 | -98.8002
3.537 | -98.619
3.5407 | -98.4294
3.5444 | -98.2747
3.5481 | -98.1753
3.5518 | -98.1313
3.5556 | -98.1202
3.5593 | -98.1202
3.563 | -98.1124
3.5667 | -98.0892
3.5704 | -98.0757
3.5742 | -98.0801
3.5779 | -98.1153
3.5816 | -98.1788
3.5853 | -98.2565
3.589 | -98.3257
3.5927 | -98.3728
3.5965 | -98.3928
3.6002 | -98.3897
3.6039 | -98.3859
3.6076 | -98.3843
3.6113 | -98.3826
3.6151 | -98.3827
3.6188 | -98.3749
3.6225 | -98.3581
3.6262 | -98.3277
3.6299 | -98.2839
3.6337 | -98.2284
3.6374 | -98.1762
3.6411 | -98.1406
3.6448 | -98.1316
3.6485 | -98.1513
3.6523 | -98.1994
3.656 | -98.2744
3.6597 | -98.3733
3.6634 | -98.5028
3.6671 | -98.6516
3.6709 | -98.7883
3.6746 | -98.9012
3.6783 | -98.9788
3.682 | -99.0161
3.6857 | -99.044
3.6894 | -99.0708
3.6932 | -99.1053
3.6969 | -99.1595
3.7006 | -99.2239
3.7043 | -99.287
3.708 | -99.336
3.7118 | -99.3509
3.7155 | -99.3269
3.7192 | -99.2622
3.7229 | -99.1529
3.7266 | -99.0155
3.7304 | -98.8593
3.7341 | -98.7042
3.7378 | -98.5819
3.7415 | -98.507
3.7452 | -98.4984
3.749 | -98.5455
3.7527 | -98.6577
3.7564 | -98.7321
3.7601 | -98.8074
3.7638 | -98.856
3.7675 | -98.8751
3.7713 | -98.8713
3.775 | -98.8524
3.7787 | -98.8295
3.7824 | -98.8344
3.7861 | -98.867
3.7899 | -98.9188
3.7936 | -98.9799
3.7973 | -99.0377
3.801 | -99.0802
3.8047 | -99.1092
3.8085 | -99.1222
3.8122 | -99.1176
3.8159 | -99.1001
3.8196 | -99.0664
3.8233 | -99.0263
3.8271 | -98.9875
3.8308 | -98.9566
3.8345 | -98.9312
3.8382 | -98.9142
3.8419 | -98.9116
3.8457 | -98.9147
3.8494 | -98.9273
3.8531 | -98.9603
3.8568 | -99.0164
3.8605 | -99.0902
3.8642 | -99.1822
3.868 | -99.2754
3.8717 | -99.3696
3.8754 | -99.4488
3.8791 | -99.4983
3.8828 | -99.5106
3.8866 | -99.4824
3.8903 | -99.4178
3.894 | -99.326
3.8977 | -99.2362
3.9014 | -99.1614
3.9052 | -99.1116
3.9089 | -99.0845
3.9126 | -99.0787
3.9163 | -99.0745
3.92 | -99.0635
3.9238 | -99.0366
3.9275 | -98.9924
3.9312 | -98.9345
3.9349 | -98.8612
3.9386 | -98.7834
3.9424 | -98.7269
3.9461 | -98.6952
3.9498 | -98.6838
3.9535 | -98.688
3.9572 | -98.7034
3.9609 | -98.7126
3.9647 | -98.7088
3.9684 | -98.692
3.9721 | -98.6856
3.9758 | -98.692
3.9795 | -98.719
3.9833 | -98.7666
3.987 | -98.8357
3.9907 | -98.9313
3.9944 | -99.0438
3.9981 | -99.1605
4.0019 | -99.286
4.0056 | -99.3854
4.0093 | -99.4559
4.013 | -99.4921
4.0167 | -99.4876
4.0205 | -99.4619
4.0242 | -99.4444
4.0279 | -99.4486
4.0316 | -99.4821
4.0353 | -99.5335
4.0391 | -99.6084
4.0428 | -99.6862
4.0465 | -99.7595
4.0502 | -99.8207
4.0539 | -99.86
4.0576 | -99.8904
4.0614 | -99.9189
4.0651 | -99.9296
4.0688 | -99.9469
4.0725 | -99.9632
4.0762 | -99.9667
4.08 | -99.9655
4.0837 | -99.9454
4.0874 | -99.9053
4.0911 | -99.8543
4.0948 | -99.7954
4.0986 | -99.7451
4.1023 | -99.7066
4.106 | -99.6882
4.1097 | -99.6911
4.1134 | -99.7164
4.1172 | -99.7661
4.1209 | -99.8293
4.1246 | -99.8955
4.1283 | -99.9559
4.132 | -100.0022
4.1358 | -100.0271
4.1395 | -100.0135
4.1432 | -99.9662
4.1469 | -99.9028
4.1506 | -99.8321
4.1543 | -99.7576
4.1581 | -99.6933
4.1618 | -99.6383
4.1655 | -99.5948
4.1692 | -99.5795
4.1729 | -99.5904
4.1767 | -99.6262
4.1804 | -99.6813
4.1841 | -99.7544
4.1878 | -99.8334
4.1915 | -99.9069
4.1953 | -99.9719
4.199 | -100.0262
4.2027 | -100.0735
4.2064 | -100.1032
4.2101 | -100.1199
4.2139 | -100.1396
4.2176 | -100.1669
4.2213 | -100.2028
4.225 | -100.2494
4.2287 | -100.2988
4.2325 | -100.3135
4.2362 | -100.2893
4.2399 | -100.2311
4.2436 | -100.1461
4.2473 | -100.0585
4.251 | -99.9772
4.2548 | -99.9001
4.2585 | -99.8295
4.2622 | -99.7624
4.2659 | -99.7019
4.2696 | -99.6272
4.2734 | -99.5486
4.2771 | -99.4661
4.2808 | -99.3994
4.2845 | -99.3608
4.2882 | -99.3562
4.292 | -99.4342
4.2957 | -99.4904
4.2994 | -99.5735
4.3031 | -99.6728
4.3068 | -99.7932
4.3106 | -99.9262
4.3143 | -100.0592
4.318 | -100.1995
4.3217 | -100.3373
4.3254 | -100.4425
4.3291 | -100.5163
4.3329 | -100.5504
4.3366 | -100.5586
4.3403 | -100.5335
4.344 | -100.4905
4.3477 | -100.4154
4.3515 | -100.327
4.3552 | -100.2473
4.3589 | -100.1831
4.3626 | -100.1372
4.3663 | -100.1194
4.3701 | -100.1234
4.3738 | -100.1357
4.3775 | -100.1418
4.3812 | -100.1451
4.3849 | -100.1445
4.3887 | -100.1327
4.3924 | -100.1167
4.3961 | -100.1099
4.3998 | -100.1119
4.4035 | -100.1299
4.4073 | -100.1649
4.411 | -100.2133
4.4147 | -100.2749
4.4184 | -100.3499
4.4221 | -100.4336
4.4258 | -100.5184
4.4296 | -100.6018
4.4333 | -100.6844
4.437 | -100.7416
4.4407 | -100.7755
4.4444 | -100.7805
4.4482 | -100.7585
4.4519 | -100.709
4.4556 | -100.6575
4.4593 | -100.6095
4.463 | -100.5701
4.4668 | -100.5393
4.4705 | -100.5253
4.4742 | -100.5264
4.4779 | -100.5429
4.4816 | -100.5691
4.4854 | -100.602
4.4891 | -100.6454
4.4928 | -100.6855
4.4965 | -100.7231
4.5002 | -100.7555
4.504 | -100.7708
4.5077 | -100.7836
4.5114 | -100.787
4.5151 | -100.7783
4.5188 | -100.7531
4.5225 | -100.7077
4.5263 | -100.633
4.53 | -100.5401
4.5337 | -100.4324
4.5374 | -100.3151
4.5411 | -100.1888
4.5449 | -100.0804
4.5486 | -99.9892
4.5523 | -99.915
4.556 | -99.8609
4.5597 | -99.8252
4.5635 | -99.8113
4.5672 | -99.8227
4.5709 | -99.856
4.5746 | -99.9051
4.5783 | -99.9503
4.5821 | -99.9824
4.5858 | -99.9961
4.5895 | -99.9971
4.5932 | -99.9922
4.5969 | -99.9816
4.6007 | -99.973
4.6044 | -99.9679
4.6081 | -99.9739
4.6118 | -99.9947
4.6155 | -100.0208
4.6192 | -100.0624
4.623 | -100.1191
4.6267 | -100.1798
4.6304 | -100.2452
4.6341 | -100.3067
4.6378 | -100.3622
4.6416 | -100.4127
4.6453 | -100.4628
4.649 | -100.5238
4.6527 | -100.5881
4.6564 | -100.64
4.6602 | -100.6743
4.6639 | -100.7031
4.6676 | -100.7304
4.6713 | -100.7507
4.675 | -100.7671
4.6788 | -100.7765
4.6825 | -100.7738
4.6862 | -100.7664
4.6899 | -100.7458
4.6936 | -100.7254
4.6974 | -100.7069
4.7011 | -100.6924
4.7048 | -100.6929
4.7085 | -100.7034
4.7122 | -100.7256
4.7159 | -100.7507
4.7197 | -100.78
4.7234 | -100.8182
4.7271 | -100.862
4.7308 | -100.8992
4.7345 | -100.9279
4.7383 | -100.9465
4.742 | -100.9439
4.7457 | -100.9229
4.7494 | -100.8765
4.7531 | -100.8195
4.7569 | -100.7609
4.7606 | -100.7143
4.7643 | -100.6822
4.768 | -100.6766
4.7717 | -100.6887
4.7755 | -100.7056
4.7792 | -100.7271
4.7829 | -100.7479
4.7866 | -100.7632
4.7903 | -100.7681
4.794 | -100.7778
4.7978 | -100.7876
4.8015 | -100.8072
4.8052 | -100.8272
4.8089 | -100.849
4.8126 | -100.8802
4.8164 | -100.9029
4.8201 | -100.9298
4.8238 | -100.9551
4.8275 | -100.9895
4.8312 | -101.0267
4.835 | -101.0712
4.8387 | -101.1169
4.8424 | -101.1631
4.8461 | -101.2044
4.8498 | -101.2387
4.8536 | -101.256
4.8573 | -101.2557
4.861 | -101.2421
4.8647 | -101.2213
4.8684 | -101.1832
4.8722 | -101.1504
4.8759 | -101.1174
4.8796 | -101.0904
4.8833 | -101.0689
4.887 | -101.0568
4.8907 | -101.0637
4.8945 | -101.0886
4.8982 | -101.1348
4.9019 | -101.1929
4.9056 | -101.2534
4.9093 | -101.3083
4.9131 | -101.3525
4.9168 | -101.3841
4.9205 | -101.3973
4.9242 | -101.3956
4.9279 | -101.3743
4.9317 | -101.346
4.9354 | -101.3113
4.9391 | -101.269
4.9428 | -101.2224
4.9465 | -101.1857
4.9503 | -101.1496
4.954 | -101.1197
4.9577 | -101.0992
4.9614 | -101.0895
4.9651 | -101.0826
4.9689 | -101.0829
4.9726 | -101.0922
4.9763 | -101.1088
4.98 | -101.1361
4.9837 | -101.1724
4.9874 | -101.2129
4.9912 | -101.2521
4.9949 | -101.2827
4.9986 | -101.3119
5.0023 | -101.3378
5.006 | -101.3522
5.0098 | -101.3557
5.0135 | -101.3415
5.0172 | -101.3078
5.0209 | -101.2658
5.0246 | -101.2154
5.0284 | -101.1649
5.0321 | -101.1115
5.0358 | -101.0535
5.0395 | -100.9931
5.0432 | -100.9322
5.047 | -100.8634
5.0507 | -100.7906
5.0544 | -100.719
5.0581 | -100.6506
5.0618 | -100.5894
5.0656 | -100.5475
5.0693 | -100.5184
5.073 | -100.5007
5.0767 | -100.4929
5.0804 | -100.4984
5.0841 | -100.5172
5.0879 | -100.5431
5.0916 | -100.5814
5.0953 | -100.6276
5.099 | -100.6708
5.1027 | -100.7183
5.1065 | -100.7716
5.1102 | -100.8355
5.1139 | -100.9074
5.1176 | -100.9818
5.1213 | -101.0604
5.1251 | -101.1408
5.1288 | -101.2175
5.1325 | -101.2891
5.1362 | -101.3483
5.1399 | -101.3994
5.1437 | -101.4322
5.1474 | -101.4624
5.1511 | -101.4807
5.1548 | -101.4827
5.1585 | -101.4662
5.1623 | -101.4445
5.166 | -101.4127
5.1697 | -101.3788
5.1734 | -101.3471
5.1771 | -101.3196
5.1808 | -101.2954
5.1846 | -101.2795
5.1883 | -101.2754
5.192 | -101.2911
5.1957 | -101.3217
5.1994 | -101.3608
5.2032 | -101.4039
5.2069 | -101.4566
5.2106 | -101.5158
5.2143 | -101.5826
5.218 | -101.6385
5.2218 | -101.688
5.2255 | -101.7256
5.2292 | -101.7631
5.2329 | -101.7877
5.2366 | -101.7908
5.2404 | -101.7757
5.2441 | -101.7481
5.2478 | -101.7055
5.2515 | -101.6535
5.2552 | -101.5964
5.2589 | -101.5388
5.2627 | -101.4779
5.2664 | -101.4189
5.2701 | -101.3685
5.2738 | -101.332
5.2775 | -101.3087
5.2813 | -101.2928
5.285 | -101.2745
5.2887 | -101.2585
5.2924 | -101.2443
5.2961 | -101.2373
5.2999 | -101.2461
5.3036 | -101.2635
5.3073 | -101.2844
5.311 | -101.3342
5.3147 | -101.4111
5.3185 | -101.4895
5.3222 | -101.5661
5.3259 | -101.6348
5.3296 | -101.6816
5.3333 | -101.6991
5.3371 | -101.7023
5.3408 | -101.6934
5.3445 | -101.6844
5.3482 | -101.6835
5.3519 | -101.699
5.3556 | -101.7249
5.3594 | -101.769
5.3631 | -101.7956
5.3668 | -101.8107
5.3705 | -101.8057
5.3742 | -101.7813
5.378 | -101.7384
5.3817 | -101.6696
5.3854 | -101.5935
5.3891 | -101.5142
5.3928 | -101.4209
5.3966 | -101.3302
5.4003 | -101.2554
5.404 | -101.1929
5.4077 | -101.148
5.4114 | -101.1265
5.4152 | -101.1306
5.4189 | -101.1552
5.4226 | -101.2022
5.4263 | -101.2604
5.43 | -101.3255
5.4338 | -101.3979
5.4375 | -101.4728
5.4412 | -101.5346
5.4449 | -101.5968
5.4486 | -101.6521
5.4523 | -101.6973
5.4561 | -101.7326
5.4598 | -101.7653
5.4635 | -101.7965
5.4672 | -101.8293
5.4709 | -101.8599
5.4747 | -101.8894
5.4784 | -101.9188
5.4821 | -101.9474
5.4858 | -101.9725
5.4895 | -101.9943
5.4933 | -102.0058
5.497 | -102.0004
5.5007 | -101.9813
5.5044 | -101.9542
5.5081 | -101.9191
5.5119 | -101.8832
5.5156 | -101.8457
5.5193 | -101.8101
5.523 | -101.7827
5.5267 | -101.7625
5.5305 | -101.749
5.5342 | -101.7455
5.5379 | -101.7523
5.5416 | -101.7629
5.5453 | -101.7784
5.549 | -101.7861
5.5528 | -101.7843
5.5565 | -101.7762
5.5602 | -101.763
5.5639 | -101.7417
5.5676 | -101.709
5.5714 | -101.6714
5.5751 | -101.6393
5.5788 | -101.613
5.5825 | -101.6026
5.5862 | -101.6046
5.59 | -101.6142
5.5937 | -101.6357
5.5974 | -101.6668
5.6011 | -101.703
5.6048 | -101.7362
5.6086 | -101.7582
5.6123 | -101.7714
5.616 | -101.7731
5.6197 | -101.7719
5.6234 | -101.7646
5.6272 | -101.7607
5.6309 | -101.7531
5.6346 | -101.7608
5.6383 | -101.7719
5.642 | -101.7825
5.6457 | -101.7928
5.6495 | -101.8039
5.6532 | -101.8218
5.6569 | -101.8456
5.6606 | -101.8667
5.6643 | -101.89
5.6681 | -101.9132
5.6718 | -101.9422
5.6755 | -101.9747
5.6792 | -102.0172
5.6829 | -102.0556
5.6867 | -102.0889
5.6904 | -102.1171
5.6941 | -102.1451
5.6978 | -102.1626
5.7015 | -102.1665
5.7053 | -102.1516
5.709 | -102.1297
5.7127 | -102.0958
5.7164 | -102.0554
5.7201 | -102.013
5.7238 | -101.9709
5.7276 | -101.9304
5.7313 | -101.9009
5.735 | -101.8859
5.7387 | -101.8722
5.7424 | -101.8623
5.7462 | -101.8566
5.7499 | -101.8555
5.7536 | -101.8551
5.7573 | -101.8512
5.761 | -101.8555
5.7648 | -101.8629
5.7685 | -101.8725
5.7722 | -101.8892
5.7759 | -101.9062
5.7796 | -101.9226
5.7834 | -101.9361
5.7871 | -101.9482
5.7908 | -101.9633
5.7945 | -101.977
5.7982 | -101.9923
5.802 | -102.0158
5.8057 | -102.0425
5.8094 | -102.0765
5.8131 | -102.1187
5.8168 | -102.1622
5.8205 | -102.2019
5.8243 | -102.2494
5.828 | -102.3015
5.8317 | -102.3422
5.8354 | -102.3606
5.8391 | -102.3689
5.8429 | -102.3609
5.8466 | -102.3308
5.8503 | -102.289
5.854 | -102.2496
5.8577 | -102.2037
5.8615 | -102.1568
5.8652 | -102.1215
5.8689 | -102.1049
5.8726 | -102.0982
5.8763 | -102.1022
5.8801 | -102.1089
5.8838 | -102.1147
5.8875 | -102.1172
5.8912 | -102.1165
5.8949 | -102.1139
5.8987 | -102.1091
5.9024 | -102.1111
5.9061 | -102.1157
5.9098 | -102.128
5.9135 | -102.1417
5.9172 | -102.1613
5.921 | -102.1905
5.9247 | -102.2183
5.9284 | -102.2349
5.9321 | -102.2457
5.9358 | -102.2532
5.9396 | -102.2485
5.9433 | -102.2376
5.947 | -102.2231
5.9507 | -102.2046
5.9544 | -102.1806
5.9582 | -102.1584
5.9619 | -102.1413
5.9656 | -102.1356
5.9693 | -102.1514
5.973 | -102.1801
5.9768 | -102.22
5.9805 | -102.2693
5.9842 | -102.318
5.9879 | -102.3634
5.9916 | -102.4086
5.9954 | -102.445
5.9991 | -102.4732
6.0028 | -102.4952
6.0065 | -102.5059
6.0102 | -102.5052
6.0139 | -102.4992
6.0177 | -102.4969
6.0214 | -102.4916
6.0251 | -102.4897
6.0288 | -102.4841
6.0325 | -102.4799
6.0363 | -102.4791
6.04 | -102.4854
6.0437 | -102.5021
6.0474 | -102.5246
6.0511 | -102.5536
6.0549 | -102.5952
6.0586 | -102.6355
6.0623 | -102.6744
6.066 | -102.7168
6.0697 | -102.7584
6.0735 | -102.799
6.0772 | -102.8396
6.0809 | -102.8709
6.0846 | -102.8932
6.0883 | -102.9057
6.0921 | -102.9012
6.0958 | -102.8893
6.0995 | -102.8739
6.1032 | -102.8554
6.1069 | -102.8351
6.1106 | -102.8071
6.1144 | -102.781
6.1181 | -102.7435
6.1218 | -102.7079
6.1255 | -102.6738
6.1292 | -102.6299
6.133 | -102.5806
6.1367 | -102.5284
6.1404 | -102.4749
6.1441 | -102.4218
6.1478 | -102.3715
6.1516 | -102.3247
6.1553 | -102.2819
6.159 | -102.2465
6.1627 | -102.2169
6.1664 | -102.1946
6.1702 | -102.1808
6.1739 | -102.1777
6.1776 | -102.183
6.1813 | -102.2019
6.185 | -102.2308
6.1887 | -102.2594
6.1925 | -102.2881
6.1962 | -102.3093
6.1999 | -102.3319
6.2036 | -102.3599
6.2073 | -102.3855
6.2111 | -102.4122
6.2148 | -102.4342
6.2185 | -102.4521
6.2222 | -102.4704
6.2259 | -102.4847
6.2297 | -102.4899
6.2334 | -102.4939
6.2371 | -102.4996
6.2408 | -102.502
6.2445 | -102.5005
6.2483 | -102.5038
6.252 | -102.5103
6.2557 | -102.5139
6.2594 | -102.5121
6.2631 | -102.5047
6.2669 | -102.4895
6.2706 | -102.4693
6.2743 | -102.4438
6.278 | -102.4177
6.2817 | -102.3936
6.2854 | -102.3782
6.2892 | -102.3728
6.2929 | -102.3723
6.2966 | -102.3724
6.3003 | -102.3822
6.304 | -102.4014
6.3078 | -102.4221
6.3115 | -102.4476
6.3152 | -102.4672
6.3189 | -102.488
6.3226 | -102.5121
6.3264 | -102.528
6.3301 | -102.5383
6.3338 | -102.5499
6.3375 | -102.5652
6.3412 | -102.5691
6.345 | -102.5608
6.3487 | -102.5469
6.3524 | -102.5323
6.3561 | -102.5107
6.3598 | -102.4881
6.3636 | -102.4687
6.3673 | -102.4523
6.371 | -102.4487
6.3747 | -102.4596
6.3784 | -102.4847
6.3821 | -102.5239
6.3859 | -102.575
6.3896 | -102.6401
6.3933 | -102.7039
6.397 | -102.7684
6.4007 | -102.8341
6.4045 | -102.8902
6.4082 | -102.9365
6.4119 | -102.9608
6.4156 | -102.9653
6.4193 | -102.9596
6.4231 | -102.9358
6.4268 | -102.8923
6.4305 | -102.8341
6.4342 | -102.7655
6.4379 | -102.7001
6.4417 | -102.6411
6.4454 | -102.589
6.4491 | -102.5419
6.4528 | -102.5086
6.4565 | -102.4856
6.4603 | -102.4713
6.464 | -102.476
6.4677 | -102.4862
6.4714 | -102.4992
6.4751 | -102.5169
6.4788 | -102.5396
6.4826 | -102.57
6.4863 | -102.6029
6.49 | -102.6396
6.4937 | -102.6721
6.4974 | -102.6991
6.5012 | -102.7255
6.5049 | -102.7551
6.5086 | -102.7817
6.5123 | -102.8043
6.516 | -102.8221
6.5198 | -102.8396
6.5235 | -102.8584
6.5272 | -102.8814
6.5309 | -102.9114
6.5346 | -102.9474
6.5384 | -102.9874
6.5421 | -103.0374
6.5458 | -103.087
6.5495 | -103.1448
6.5532 | -103.2032
6.5569 | -103.2573
6.5607 | -103.3106
6.5644 | -103.3547
6.5681 | -103.3946
6.5718 | -103.4273
6.5755 | -103.4411
6.5793 | -103.4338
6.583 | -103.4033
6.5867 | -103.3552
6.5904 | -103.2959
6.5941 | -103.2233
6.5979 | -103.146
6.6016 | -103.0642
6.6053 | -102.9837
6.609 | -102.91
6.6127 | -102.8483
6.6165 | -102.7973
6.6202 | -102.7591
6.6239 | -102.7391
6.6276 | -102.7329
6.6313 | -102.7404
6.6351 | -102.7641
6.6388 | -102.8069
6.6425 | -102.8586
6.6462 | -102.9196
6.6499 | -102.9782
6.6536 | -103.0394
6.6574 | -103.0985
6.6611 | -103.1457
6.6648 | -103.181
6.6685 | -103.1946
6.6722 | -103.1915
6.676 | -103.1883
6.6797 | -103.1811
6.6834 | -103.1603
6.6871 | -103.1328
6.6908 | -103.1038
6.6946 | -103.0763
6.6983 | -103.0432
6.702 | -103.0142
6.7057 | -102.9855
6.7094 | -102.9636
6.7132 | -102.9453
6.7169 | -102.9365
6.7206 | -102.9349
6.7243 | -102.9388
6.728 | -102.9515
6.7318 | -102.9724
6.7355 | -102.9982
6.7392 | -103.0278
6.7429 | -103.0649
6.7466 | -103.1069
6.7503 | -103.1516
6.7541 | -103.1959
6.7578 | -103.2383
6.7615 | -103.2758
6.7652 | -103.3114
6.7689 | -103.3486
6.7727 | -103.3802
6.7764 | -103.3946
6.7801 | -103.4037
6.7838 | -103.416
6.7875 | -103.4246
6.7913 | -103.4256
6.795 | -103.4218
6.7987 | -103.4171
6.8024 | -103.4142
6.8061 | -103.413
6.8099 | -103.4115
6.8136 | -103.4086
6.8173 | -103.4036
6.821 | -103.3973
6.8247 | -103.3903
6.8285 | -103.3765
6.8322 | -103.3567
6.8359 | -103.3317
6.8396 | -103.3094
6.8433 | -103.286
6.847 | -103.2592
6.8508 | -103.2271
6.8545 | -103.1932
6.8582 | -103.1641
6.8619 | -103.1352
6.8656 | -103.0977
6.8694 | -103.0681
6.8731 | -103.041
6.8768 | -103.0238
6.8805 | -103.0121
6.8842 | -102.9974
6.888 | -102.9848
6.8917 | -102.9756
6.8954 | -102.9744
6.8991 | -102.9741
6.9028 | -102.9712
6.9066 | -102.9745
6.9103 | -102.9891
6.914 | -103.0124
6.9177 | -103.0415
6.9214 | -103.0804
6.9252 | -103.127
6.9289 | -103.1759
6.9326 | -103.2322
6.9363 | -103.2903
6.94 | -103.3418
6.9437 | -103.3914
6.9475 | -103.4359
6.9512 | -103.4714
6.9549 | -103.4915
6.9586 | -103.4983
6.9623 | -103.4942
6.9661 | -103.4833
6.9698 | -103.462
6.9735 | -103.4418
6.9772 | -103.4273
6.9809 | -103.4113
6.9847 | -103.3877
6.9884 | -103.3811
6.9921 | -103.3788
6.9958 | -103.3799
6.9995 | -103.3763
7.0033 | -103.3728
7.007 | -103.3682
7.0107 | -103.3655
7.0144 | -103.3534
7.0181 | -103.3385
7.0218 | -103.3231
7.0256 | -103.2999
7.0293 | -103.2757
7.033 | -103.2566
7.0367 | -103.2308
7.0404 | -103.1994
7.0442 | -103.1625
7.0479 | -103.1295
7.0516 | -103.0978
7.0553 | -103.0661
7.059 | -103.0369
7.0628 | -103.0081
7.0665 | -102.9805
7.0702 | -102.9597
7.0739 | -102.9445
7.0776 | -102.9368
7.0814 | -102.9375
7.0851 | -102.9474
7.0888 | -102.9719
7.0925 | -103.015
7.0962 | -103.0613
7.1 | -103.1176
7.1037 | -103.1805
7.1074 | -103.2565
7.1111 | -103.3391
7.1148 | -103.4255
7.1185 | -103.5076
7.1223 | -103.5895
7.126 | -103.6677
7.1297 | -103.7299
7.1334 | -103.7784
7.1371 | -103.8018
7.1409 | -103.8059
7.1446 | -103.7983
7.1483 | -103.7711
7.152 | -103.7294
7.1557 | -103.6728
7.1595 | -103.6059
7.1632 | -103.5393
7.1669 | -103.4713
7.1706 | -103.4067
7.1743 | -103.3516
7.1781 | -103.3062
7.1818 | -103.2656
7.1855 | -103.2336
7.1892 | -103.2125
7.1929 | -103.1992
7.1967 | -103.1944
7.2004 | -103.1991
7.2041 | -103.2093
7.2078 | -103.2233
7.2115 | -103.2411
7.2152 | -103.2546
7.219 | -103.2691
7.2227 | -103.2865
7.2264 | -103.3045
7.2301 | -103.3231
7.2338 | -103.3418
7.2376 | -103.3659
7.2413 | -103.3915
7.245 | -103.4207
7.2487 | -103.4494
7.2524 | -103.4799
7.2562 | -103.5153
7.2599 | -103.5453
7.2636 | -103.564
7.2673 | -103.5798
7.271 | -103.5853
7.2748 | -103.5824
7.2785 | -103.5681
7.2822 | -103.5457
7.2859 | -103.5118
7.2896 | -103.4737
7.2934 | -103.4356
7.2971 | -103.3929
7.3008 | -103.3496
7.3045 | -103.3082
7.3082 | -103.2737
7.3119 | -103.2502
7.3157 | -103.2397
7.3194 | -103.2415
7.3231 | -103.2532
7.3268 | -103.2721
7.3305 | -103.2948
7.3343 | -103.3292
7.338 | -103.3776
7.3417 | -103.4307
7.3454 | -103.4799
7.3491 | -103.5278
7.3529 | -103.5745
7.3566 | -103.614
7.3603 | -103.6446
7.364 | -103.6702
7.3677 | -103.6892
7.3715 | -103.6958
7.3752 | -103.6919
7.3789 | -103.6807
7.3826 | -103.6601
7.3863 | -103.6346
7.3901 | -103.6001
7.3938 | -103.5696
7.3975 | -103.5392
7.4012 | -103.5103
7.4049 | -103.4852
7.4086 | -103.4611
7.4124 | -103.4413
7.4161 | -103.4157
7.4198 | -103.3922
7.4235 | -103.3752
7.4272 | -103.3612
7.431 | -103.3546
7.4347 | -103.344
7.4384 | -103.3333
7.4421 | -103.3233
7.4458 | -103.3135
7.4496 | -103.3104
7.4533 | -103.3184
7.457 | -103.3359
7.4607 | -103.3631
7.4644 | -103.4009
7.4682 | -103.4488
7.4719 | -103.5059
7.4756 | -103.5713
7.4793 | -103.6503
7.483 | -103.7357
7.4867 | -103.8249
7.4905 | -103.9188
7.4942 | -104.0091
7.4979 | -104.0955
7.5016 | -104.1679
7.5053 | -104.2324
7.5091 | -104.2457
7.5128 | -104.2786
7.5165 | -104.2981
7.5202 | -104.3066
7.5239 | -104.3008
7.5277 | -104.2868
7.5314 | -104.2585
7.5351 | -104.2131
7.5388 | -104.1654
7.5425 | -104.1149
7.5463 | -104.0673
7.55 | -104.0175
7.5537 | -103.9637
7.5574 | -103.9137
7.5611 | -103.8647
7.5649 | -103.8178
7.5686 | -103.7706
7.5723 | -103.7263
7.576 | -103.6876
7.5797 | -103.6518
7.5835 | -103.6239
7.5872 | -103.6024
7.5909 | -103.579
7.5946 | -103.5593
7.5983 | -103.5444
7.602 | -103.5345
7.6058 | -103.5322
7.6095 | -103.5347
7.6132 | -103.5439
7.6169 | -103.5569
7.6206 | -103.5725
7.6244 | -103.5912
7.6281 | -103.6171
7.6318 | -103.6477
7.6355 | -103.6778
7.6392 | -103.7073
7.643 | -103.7397
7.6467 | -103.7664
7.6504 | -103.7958
7.6541 | -103.8204
7.6578 | -103.8396
7.6616 | -103.8565
7.6653 | -103.8719
7.669 | -103.8809
7.6727 | -103.8857
7.6764 | -103.8943
7.6801 | -103.9016
7.6839 | -103.9069
7.6876 | -103.9076
7.6913 | -103.9043
7.695 | -103.8979
7.6987 | -103.8894
7.7025 | -103.8783
7.7062 | -103.8692
7.7099 | -103.8616
7.7136 | -103.8553
7.7173 | -103.8454
7.7211 | -103.8367
7.7248 | -103.8256
7.7285 | -103.8084
7.7322 | -103.7931
7.7359 | -103.7837
7.7397 | -103.7767
7.7434 | -103.7603
7.7471 | -103.7413
7.7508 | -103.7321
7.7545 | -103.7261
7.7583 | -103.7215
7.762 | -103.7228
7.7657 | -103.7241
7.7694 | -103.7321
7.7731 | -103.748
7.7768 | -103.7691
7.7806 | -103.797
7.7843 | -103.839
7.788 | -103.882
7.7917 | -103.9254
7.7954 | -103.9737
7.7992 | -104.0271
7.8029 | -104.0804
7.8066 | -104.1282
7.8103 | -104.1736
7.814 | -104.2115
7.8178 | -104.2453
7.8215 | -104.2671
7.8252 | -104.2728
7.8289 | -104.2723
7.8326 | -104.2718
7.8364 | -104.265
7.8401 | -104.2506
7.8438 | -104.234
7.8475 | -104.2138
7.8512 | -104.1878
7.855 | -104.1615
7.8587 | -104.1425
7.8624 | -104.1251
7.8661 | -104.1008
7.8698 | -104.0802
7.8735 | -104.0648
7.8773 | -104.0546
7.881 | -104.0463
7.8847 | -104.0411
7.8884 | -104.0399
7.8921 | -104.0467
7.8959 | -104.0537
7.8996 | -104.0592
7.9033 | -104.0692
7.907 | -104.0857
7.9107 | -104.1048
7.9145 | -104.1201
7.9182 | -104.1372
7.9219 | -104.1569
7.9256 | -104.1654
7.9293 | -104.1706
7.9331 | -104.1707
7.9368 | -104.1666
7.9405 | -104.1618
7.9442 | -104.1514
7.9479 | -104.1358
7.9516 | -104.1162
7.9554 | -104.1017
7.9591 | -104.0906
7.9628 | -104.0729
7.9665 | -104.0519
7.9702 | -104.0352
7.974 | -104.0258
7.9777 | -104.0169
7.9814 | -104.0145
7.9851 | -104.0169
7.9888 | -104.015
7.9926 | -104.016
7.9963 | -104.025
8 | -104.0339
8.0037 | -104.0425
8.0074 | -104.0517
8.0112 | -104.0738
8.0149 | -104.0962
8.0186 | -104.1165
8.0223 | -104.1376
8.026 | -104.1577
8.0298 | -104.1789
8.0335 | -104.1998
8.0372 | -104.2182
8.0409 | -104.2341
8.0446 | -104.2444
8.0483 | -104.2586
8.0521 | -104.2733
8.0558 | -104.2906
8.0595 | -104.3081
8.0632 | -104.3212
8.0669 | -104.3334
8.0707 | -104.3452
8.0744 | -104.3543
8.0781 | -104.3636
8.0818 | -104.3786
8.0855 | -104.3967
8.0893 | -104.4035
8.093 | -104.4047
8.0967 | -104.4041
8.1004 | -104.3984
8.1041 | -104.3911
8.1079 | -104.3848
8.1116 | -104.3769
8.1153 | -104.3611
8.119 | -104.341
8.1227 | -104.3263
8.1265 | -104.3131
8.1302 | -104.2987
8.1339 | -104.2908
8.1376 | -104.2923
8.1413 | -104.2933
8.145 | -104.2971
8.1488 | -104.3085
8.1525 | -104.3217
8.1562 | -104.3401
8.1599 | -104.3649
8.1636 | -104.3934
8.1674 | -104.4151
8.1711 | -104.4351
8.1748 | -104.4547
8.1785 | -104.4709
8.1822 | -104.4855
8.186 | -104.4986
8.1897 | -104.5092
8.1934 | -104.5182
8.1971 | -104.5178
8.2008 | -104.5103
8.2046 | -104.4999
8.2083 | -104.4852
8.212 | -104.4749
8.2157 | -104.4621
8.2194 | -104.4551
8.2232 | -104.4504
8.2269 | -104.4448
8.2306 | -104.4412
8.2343 | -104.4407
8.238 | -104.4464
8.2417 | -104.4599
8.2455 | -104.4698
8.2492 | -104.4765
8.2529 | -104.4826
8.2566 | -104.4886
8.2603 | -104.4929
8.2641 | -104.4966
8.2678 | -104.4972
8.2715 | -104.4963
8.2752 | -104.4953
8.2789 | -104.4846
8.2827 | -104.4684
8.2864 | -104.4532
8.2901 | -104.4387
8.2938 | -104.4193
8.2975 | -104.4025
8.3013 | -104.3822
8.305 | -104.3569
8.3087 | -104.3346
8.3124 | -104.3108
8.3161 | -104.287
8.3199 | -104.2665
8.3236 | -104.2476
8.3273 | -104.2325
8.331 | -104.2171
8.3347 | -104.2011
8.3384 | -104.1918
8.3422 | -104.186
8.3459 | -104.1843
8.3496 | -104.1917
8.3533 | -104.2022
8.357 | -104.2095
8.3608 | -104.2181
8.3645 | -104.2263
8.3682 | -104.2359
8.3719 | -104.247
8.3756 | -104.2533
8.3794 | -104.2586
8.3831 | -104.2608
8.3868 | -104.2571
8.3905 | -104.251
8.3942 | -104.2417
8.398 | -104.2298
8.4017 | -104.2204
8.4054 | -104.215
8.4091 | -104.2082
8.4128 | -104.2058
8.4166 | -104.2068
8.4203 | -104.2079
8.424 | -104.22
8.4277 | -104.2403
8.4314 | -104.2649
8.4351 | -104.2931
8.4389 | -104.3297
8.4426 | -104.374
8.4463 | -104.4197
8.45 | -104.4659
8.4537 | -104.5169
8.4575 | -104.5662
8.4612 | -104.6082
8.4649 | -104.645
8.4686 | -104.6785
8.4723 | -104.705
8.4761 | -104.7282
8.4798 | -104.7529
8.4835 | -104.7754
8.4872 | -104.7942
8.4909 | -104.811
8.4947 | -104.8191
8.4984 | -104.8318
8.5021 | -104.8431
8.5058 | -104.8475
8.5095 | -104.8542
8.5132 | -104.8623
8.517 | -104.8679
8.5207 | -104.8699
8.5244 | -104.8744
8.5281 | -104.8748
8.5318 | -104.873
8.5356 | -104.8713
8.5393 | -104.8703
8.543 | -104.8653
8.5467 | -104.8551
8.5504 | -104.8357
8.5542 | -104.8176
8.5579 | -104.8014
8.5616 | -104.7749
8.5653 | -104.7416
8.569 | -104.7067
8.5728 | -104.6731
8.5765 | -104.6392
8.5802 | -104.606
8.5839 | -104.5782
8.5876 | -104.5504
8.5914 | -104.5201
8.5951 | -104.4914
8.5988 | -104.4687
8.6025 | -104.4512
8.6062 | -104.4339
8.6099 | -104.416
8.6137 | -104.4054
8.6174 | -104.3951
8.6211 | -104.3914
8.6248 | -104.3864
8.6285 | -104.3841
8.6323 | -104.386
8.636 | -104.3869
8.6397 | -104.3864
8.6434 | -104.385
8.6471 | -104.3877
8.6509 | -104.3892
8.6546 | -104.3922
8.6583 | -104.3981
8.662 | -104.4052
8.6657 | -104.4116
8.6695 | -104.4234
8.6732 | -104.4374
8.6769 | -104.4577
8.6806 | -104.4765
8.6843 | -104.4933
8.6881 | -104.5153
8.6918 | -104.5407
8.6955 | -104.5604
8.6992 | -104.5707
8.7029 | -104.5739
8.7066 | -104.5696
8.7104 | -104.5623
8.7141 | -104.5507
8.7178 | -104.5361
8.7215 | -104.5192
8.7252 | -104.5002
8.729 | -104.4757
8.7327 | -104.4491
8.7364 | -104.4212
8.7401 | -104.3974
8.7438 | -104.374
8.7476 | -104.358
8.7513 | -104.3453
8.755 | -104.3338
8.7587 | -104.3265
8.7624 | -104.3275
8.7662 | -104.3313
8.7699 | -104.3428
8.7736 | -104.3602
8.7773 | -104.3779
8.781 | -104.3988
8.7848 | -104.4232
8.7885 | -104.4438
8.7922 | -104.4676
8.7959 | -104.4947
8.7996 | -104.5171
8.8033 | -104.5343
8.8071 | -104.5493
8.8108 | -104.5646
8.8145 | -104.5803
8.8182 | -104.5923
8.8219 | -104.5972
8.8257 | -104.6032
8.8294 | -104.61
8.8331 | -104.6178
8.8368 | -104.6259
8.8405 | -104.6335
8.8443 | -104.6377
8.848 | -104.6432
8.8517 | -104.6562
8.8554 | -104.6689
8.8591 | -104.6793
8.8629 | -104.6913
8.8666 | -104.706
8.8703 | -104.7173
8.874 | -104.724
8.8777 | -104.726
8.8814 | -104.726
8.8852 | -104.7282
8.8889 | -104.7352
8.8926 | -104.7373
8.8963 | -104.7448
8.9 | -104.7551
8.9038 | -104.7699
8.9075 | -104.7889
8.9112 | -104.8094
8.9149 | -104.8324
8.9186 | -104.8629
8.9224 | -104.8862
8.9261 | -104.9003
8.9298 | -104.9117
8.9335 | -104.9238
8.9372 | -104.9294
8.941 | -104.9301
8.9447 | -104.9233
8.9484 | -104.9104
8.9521 | -104.8989
8.9558 | -104.8829
8.9596 | -104.8603
8.9633 | -104.8323
8.967 | -104.7994
8.9707 | -104.7652
8.9744 | -104.7319
8.9781 | -104.7017
8.9819 | -104.6696
8.9856 | -104.6407
8.9893 | -104.6219
8.993 | -104.6053
8.9967 | -104.5942
9.0005 | -104.5893
9.0042 | -104.5941
9.0079 | -104.6094
9.0116 | -104.63
9.0153 | -104.6493
9.0191 | -104.6692
9.0228 | -104.6962
9.0265 | -104.7277
9.0302 | -104.7581
9.0339 | -104.7879
9.0377 | -104.8199
9.0414 | -104.8443
9.0451 | -104.8599
9.0488 | -104.8656
9.0525 | -104.8724
9.0563 | -104.8769
9.06 | -104.8747
9.0637 | -104.8633
9.0674 | -104.8478
9.0711 | -104.8294
9.0748 | -104.8148
9.0786 | -104.798
9.0823 | -104.779
9.086 | -104.7632
9.0897 | -104.7478
9.0934 | -104.7367
9.0972 | -104.7344
9.1009 | -104.7348
9.1046 | -104.7396
9.1083 | -104.747
9.112 | -104.759
9.1158 | -104.7789
9.1195 | -104.8012
9.1232 | -104.8177
9.1269 | -104.8385
9.1306 | -104.8612
9.1344 | -104.885
9.1381 | -104.9031
9.1418 | -104.9187
9.1455 | -104.9344
9.1492 | -104.9529
9.153 | -104.9678
9.1567 | -104.9812
9.1604 | -104.9907
9.1641 | -105.0042
9.1678 | -105.0181
9.1715 | -105.0344
9.1753 | -105.0577
9.179 | -105.0838
9.1827 | -105.111
9.1864 | -105.1331
9.1901 | -105.1583
9.1939 | -105.1829
9.1976 | -105.2073
9.2013 | -105.2297
9.205 | -105.2478
9.2087 | -105.2649
9.2125 | -105.2807
9.2162 | -105.2947
9.2199 | -105.3079
9.2236 | -105.3222
9.2273 | -105.3308
9.2311 | -105.3291
9.2348 | -105.3268
9.2385 | -105.3229
9.2422 | -105.313
9.2459 | -105.2969
9.2497 | -105.2793
9.2534 | -105.2598
9.2571 | -105.2352
9.2608 | -105.2097
9.2645 | -105.1811
9.2682 | -105.1467
9.272 | -105.1077
9.2757 | -105.0695
9.2794 | -105.0341
9.2831 | -104.9983
9.2868 | -104.9624
9.2906 | -104.9291
9.2943 | -104.8972
9.298 | -104.8685
9.3017 | -104.8421
9.3054 | -104.8229
9.3092 | -104.8075
9.3129 | -104.7968
9.3166 | -104.793
9.3203 | -104.7954
9.324 | -104.8093
9.3278 | -104.8244
9.3315 | -104.8416
9.3352 | -104.8606
9.3389 | -104.8855
9.3426 | -104.9181
9.3464 | -104.9474
9.3501 | -104.9786
9.3538 | -105.0106
9.3575 | -105.043
9.3612 | -105.072
9.3649 | -105.1005
9.3687 | -105.1268
9.3724 | -105.1446
9.3761 | -105.1613
9.3798 | -105.1702
9.3835 | -105.1769
9.3873 | -105.1759
9.391 | -105.1728
9.3947 | -105.1656
9.3984 | -105.1532
9.4021 | -105.1382
9.4059 | -105.1266
9.4096 | -105.1192
9.4133 | -105.1134
9.417 | -105.1065
9.4207 | -105.1042
9.4245 | -105.1044
9.4282 | -105.1087
9.4319 | -105.117
9.4356 | -105.1318
9.4393 | -105.1514
9.443 | -105.1745
9.4468 | -105.1993
9.4505 | -105.2319
9.4542 | -105.267
9.4579 | -105.3039
9.4616 | -105.333
9.4654 | -105.3619
9.4691 | -105.393
9.4728 | -105.4237
9.4765 | -105.4508
9.4802 | -105.4667
9.484 | -105.4786
9.4877 | -105.4867
9.4914 | -105.4819
9.4951 | -105.4716
9.4988 | -105.4621
9.5026 | -105.4449
9.5063 | -105.4228
9.51 | -105.396
9.5137 | -105.3649
9.5174 | -105.3356
9.5212 | -105.2989
9.5249 | -105.2592
9.5286 | -105.2165
9.5323 | -105.1688
9.536 | -105.1177
9.5397 | -105.0686
9.5435 | -105.0208
9.5472 | -104.9711
9.5509 | -104.9163
9.5546 | -104.8622
9.5583 | -104.815
9.5621 | -104.7719
9.5658 | -104.7305
9.5695 | -104.6908
9.5732 | -104.6563
9.5769 | -104.6302
9.5807 | -104.6076
9.5844 | -104.5932
9.5881 | -104.584
9.5918 | -104.5835
9.5955 | -104.5902
9.5993 | -104.6101
9.603 | -104.6372
9.6067 | -104.6699
9.6104 | -104.7057
9.6141 | -104.7444
9.6179 | -104.7939
9.6216 | -104.8493
9.6253 | -104.9027
9.629 | -104.957
9.6327 | -105.0082
9.6364 | -105.0596
9.6402 | -105.1103
9.6439 | -105.1545
9.6476 | -105.1974
9.6513 | -105.2325
9.655 | -105.2398
9.6588 | -105.2643
9.6625 | -105.2865
9.6662 | -105.299
9.6699 | -105.3082
9.6736 | -105.3204
9.6774 | -105.3323
9.6811 | -105.33
9.6848 | -105.3223
9.6885 | -105.3169
9.6922 | -105.3125
9.696 | -105.3091
9.6997 | -105.3012
9.7034 | -105.2983
9.7071 | -105.2981
9.7108 | -105.2953
9.7146 | -105.291
9.7183 | -105.2853
9.722 | -105.2822
9.7257 | -105.2808
9.7294 | -105.2789
9.7331 | -105.2779
9.7369 | -105.2765
9.7406 | -105.2723
9.7443 | -105.2694
9.748 | -105.2659
9.7517 | -105.2609
9.7555 | -105.2538
9.7592 | -105.2447
9.7629 | -105.2369
9.7666 | -105.2327
9.7703 | -105.2305
9.7741 | -105.2278
9.7778 | -105.2224
9.7815 | -105.2175
9.7852 | -105.2163
9.7889 | -105.2163
9.7927 | -105.217
9.7964 | -105.2198
9.8001 | -105.2284
9.8038 | -105.2322
9.8075 | -105.2384
9.8113 | -105.2476
9.815 | -105.2563
9.8187 | -105.2647
9.8224 | -105.2739
9.8261 | -105.2876
9.8298 | -105.2975
9.8336 | -105.3066
9.8373 | -105.3148
9.841 | -105.3174
9.8447 | -105.3208
9.8484 | -105.3206
9.8522 | -105.3221
9.8559 | -105.3305
9.8596 | -105.3368
9.8633 | -105.3439
9.867 | -105.353
9.8708 | -105.3601
9.8745 | -105.3636
9.8782 | -105.3681
9.8819 | -105.3729
9.8856 | -105.373
9.8894 | -105.3729
9.8931 | -105.3725
9.8968 | -105.3709
9.9005 | -105.3689
9.9042 | -105.363
9.9079 | -105.3545
9.9117 | -105.3437
9.9154 | -105.332
9.9191 | -105.3198
9.9228 | -105.3072
9.9265 | -105.2991
9.9303 | -105.2903
9.934 | -105.2807
9.9377 | -105.2718
9.9414 | -105.2645
9.9451 | -105.2572
9.9489 | -105.2501
9.9526 | -105.2401
9.9563 | -105.2349
9.96 | -105.2304
9.9637 | -105.224
9.9675 | -105.2147
9.9712 | -105.2086
9.9749 | -105.2047
9.9786 | -105.2
9.9823 | -105.1952
9.9861 | -105.1877
9.9898 | -105.1814
9.9935 | -105.1739
9.9972 | -105.1624
10.0009 | -105.1474
10.0046 | -105.1338
10.0084 | -105.1225
10.0121 | -105.1118
10.0158 | -105.1031
10.0195 | -105.1002
10.0232 | -105.0984
10.027 | -105.0972
10.0307 | -105.0982
10.0344 | -105.1001
10.0381 | -105.103
10.0418 | -105.1067
10.0456 | -105.1105
10.0493 | -105.1175
10.053 | -105.1298
10.0567 | -105.1458
10.0604 | -105.1619
10.0642 | -105.1803
10.0679 | -105.1998
10.0716 | -105.2185
10.0753 | -105.2388
10.079 | -105.2555
10.0828 | -105.2751
10.0865 | -105.2969
10.0902 | -105.3175
10.0939 | -105.3418
10.0976 | -105.3677
10.1013 | -105.3839
10.1051 | -105.3927
10.1088 | -105.4017
10.1125 | -105.4075
10.1162 | -105.408
10.1199 | -105.4032
10.1237 | -105.3942
10.1274 | -105.388
10.1311 | -105.3824
10.1348 | -105.3706
10.1385 | -105.3585
10.1423 | -105.3474
10.146 | -105.3365
10.1497 | -105.3235
10.1534 | -105.3082
10.1571 | -105.298
10.1609 | -105.2906
10.1646 | -105.2849
10.1683 | -105.284
10.172 | -105.2903
10.1757 | -105.3007
10.1795 | -105.3083
10.1832 | -105.3211
10.1869 | -105.3402
10.1906 | -105.3592
10.1943 | -105.3829
10.198 | -105.3995
10.2018 | -105.4147
10.2055 | -105.4319
10.2092 | -105.4494
10.2129 | -105.4647
10.2166 | -105.4825
10.2204 | -105.4985
10.2241 | -105.5159
10.2278 | -105.5346
10.2315 | -105.5481
10.2352 | -105.5573
10.239 | -105.5658
10.2427 | -105.5737
10.2464 | -105.5754
10.2501 | -105.5785
10.2538 | -105.5808
10.2576 | -105.5799
10.2613 | -105.5787
10.265 | -105.5778
10.2687 | -105.5757
10.2724 | -105.5701
10.2762 | -105.5626
10.2799 | -105.5547
10.2836 | -105.5467
10.2873 | -105.542
10.291 | -105.5378
10.2947 | -105.5337
10.2985 | -105.5267
10.3022 | -105.5202
10.3059 | -105.513
10.3096 | -105.5048
10.3133 | -105.4982
10.3171 | -105.4911
10.3208 | -105.4852
10.3245 | -105.4811
10.3282 | -105.4829
10.3319 | -105.4856
10.3357 | -105.4881
10.3394 | -105.4972
10.3431 | -105.5099
10.3468 | -105.5275
10.3505 | -105.5463
10.3543 | -105.562
10.358 | -105.5776
10.3617 | -105.5961
10.3654 | -105.6154
10.3691 | -105.6338
10.3728 | -105.6509
10.3766 | -105.6603
10.3803 | -105.6679
10.384 | -105.6777
10.3877 | -105.6855
10.3914 | -105.6877
10.3952 | -105.6879
10.3989 | -105.6908
10.4026 | -105.6887
10.4063 | -105.6879
10.41 | -105.6847
10.4138 | -105.6829
10.4175 | -105.6784
10.4212 | -105.6748
10.4249 | -105.6717
10.4286 | -105.6667
10.4324 | -105.663
10.4361 | -105.6645
10.4398 | -105.6654
10.4435 | -105.6632
10.4472 | -105.6597
10.451 | -105.6571
10.4547 | -105.6521
10.4584 | -105.6472
10.4621 | -105.646
10.4658 | -105.6398
10.4695 | -105.633
10.4733 | -105.6245
10.477 | -105.6181
10.4807 | -105.608
10.4844 | -105.5965
10.4881 | -105.5813
10.4919 | -105.5627
10.4956 | -105.547
10.4993 | -105.5327
10.503 | -105.5205
10.5067 | -105.5107
10.5105 | -105.502
10.5142 | -105.4929
10.5179 | -105.4841
10.5216 | -105.4719
10.5253 | -105.4642
10.5291 | -105.463
10.5328 | -105.4641
10.5365 | -105.4652
10.5402 | -105.4706
10.5439 | -105.4761
10.5477 | -105.4845
10.5514 | -105.4999
10.5551 | -105.5207
10.5588 | -105.5409
10.5625 | -105.5604
10.5662 | -105.5799
10.57 | -105.5999
10.5737 | -105.6223
10.5774 | -105.65
10.5811 | -105.6786
10.5848 | -105.7057
10.5886 | -105.7274
10.5923 | -105.7452
10.596 | -105.7695
10.5997 | -105.7908
10.6034 | -105.8097
10.6072 | -105.8228
10.6109 | -105.824
10.6146 | -105.8183
10.6183 | -105.8158
10.622 | -105.8102
10.6258 | -105.7988
10.6295 | -105.7798
10.6332 | -105.758
10.6369 | -105.7342
10.6406 | -105.7134
10.6444 | -105.6894
10.6481 | -105.6608
10.6518 | -105.6321
10.6555 | -105.6028
10.6592 | -105.5779
10.6629 | -105.5524
10.6667 | -105.5275
10.6704 | -105.5057
10.6741 | -105.4894
10.6778 | -105.4743
10.6815 | -105.4565
10.6853 | -105.4389
10.689 | -105.4257
10.6927 | -105.4139
10.6964 | -105.4064
10.7001 | -105.3958
10.7039 | -105.3803
10.7076 | -105.3673
10.7113 | -105.3604
10.715 | -105.3533
10.7187 | -105.3468
10.7225 | -105.3434
10.7262 | -105.3395
10.7299 | -105.3375
10.7336 | -105.334
10.7373 | -105.336
10.7411 | -105.3428
10.7448 | -105.3517
10.7485 | -105.3642
10.7522 | -105.3769
10.7559 | -105.3949
10.7596 | -105.4176
10.7634 | -105.4437
10.7671 | -105.4767
10.7708 | -105.5089
10.7745 | -105.5394
10.7782 | -105.5685
10.782 | -105.5955
10.7857 | -105.6245
10.7894 | -105.6572
10.7931 | -105.6876
10.7968 | -105.7137
10.8006 | -105.7363
10.8043 | -105.7536
10.808 | -105.7656
10.8117 | -105.779
10.8154 | -105.7883
10.8192 | -105.7915
10.8229 | -105.7897
10.8266 | -105.7839
10.8303 | -105.7727
10.834 | -105.7575
10.8377 | -105.7449
10.8415 | -105.7319
10.8452 | -105.7154
10.8489 | -105.6952
10.8526 | -105.6711
10.8563 | -105.6496
10.8601 | -105.6324
10.8638 | -105.6148
10.8675 | -105.5936
10.8712 | -105.5751
10.8749 | -105.5613
10.8787 | -105.5505
10.8824 | -105.5399
10.8861 | -105.5325
10.8898 | -105.5271
10.8935 | -105.5217
10.8973 | -105.5182
10.901 | -105.5156
10.9047 | -105.5153
10.9084 | -105.5163
10.9121 | -105.5195
10.9159 | -105.5219
10.9196 | -105.5219
10.9233 | -105.5235
10.927 | -105.5222
10.9307 | -105.5256
10.9344 | -105.5275
10.9382 | -105.528
10.9419 | -105.5284
10.9456 | -105.5255
10.9493 | -105.517
10.953 | -105.5083
10.9568 | -105.5006
10.9605 | -105.4948
10.9642 | -105.4888
10.9679 | -105.4792
10.9716 | -105.4671
10.9754 | -105.4577
10.9791 | -105.4528
10.9828 | -105.447
10.9865 | -105.4394
10.9902 | -105.4348
10.994 | -105.4343
10.9977 | -105.4356
11.0014 | -105.44
11.0051 | -105.4462
11.0088 | -105.4573
11.0126 | -105.4722
11.0163 | -105.486
11.02 | -105.4972
11.0237 | -105.5098
11.0274 | -105.5247
11.0311 | -105.5404
11.0349 | -105.5546
11.0386 | -105.565
11.0423 | -105.5759
11.046 | -105.5856
11.0497 | -105.5911
11.0535 | -105.5922
11.0572 | -105.585
11.0609 | -105.5788
11.0646 | -105.5711
11.0683 | -105.5588
11.0721 | -105.54
11.0758 | -105.5213
11.0795 | -105.5013
11.0832 | -105.4809
11.0869 | -105.4601
11.0907 | -105.4348
11.0944 | -105.4121
11.0981 | -105.3921
11.1018 | -105.379
11.1055 | -105.3681
11.1093 | -105.3604
11.113 | -105.3567
11.1167 | -105.355
11.1204 | -105.3557
11.1241 | -105.3628
11.1278 | -105.3741
11.1316 | -105.3892
11.1353 | -105.4064
11.139 | -105.4267
11.1427 | -105.4543
11.1464 | -105.4855
11.1502 | -105.5149
11.1539 | -105.5429
11.1576 | -105.5713
11.1613 | -105.6022
11.165 | -105.6358
11.1688 | -105.6704
11.1725 | -105.7039
11.1762 | -105.7338
11.1799 | -105.7625
11.1836 | -105.7862
11.1874 | -105.8062
11.1911 | -105.824
11.1948 | -105.8349
11.1985 | -105.8435
11.2022 | -105.846
11.206 | -105.8461
11.2097 | -105.8451
11.2134 | -105.8398
11.2171 | -105.8355
11.2208 | -105.8241
11.2245 | -105.8113
11.2283 | -105.7982
11.232 | -105.7872
11.2357 | -105.777
11.2394 | -105.7654
11.2431 | -105.7494
11.2469 | -105.7328
11.2506 | -105.718
11.2543 | -105.7014
11.258 | -105.6796
11.2617 | -105.6548
11.2655 | -105.6358
11.2692 | -105.6147
11.2729 | -105.5937
11.2766 | -105.5738
11.2803 | -105.5582
11.2841 | -105.5451
11.2878 | -105.5323
11.2915 | -105.5223
11.2952 | -105.5077
11.2989 | -105.4916
11.3026 | -105.4799
11.3064 | -105.4713
11.3101 | -105.4647
11.3138 | -105.4621
11.3175 | -105.4627
11.3212 | -105.4622
11.325 | -105.4631
11.3287 | -105.4689
11.3324 | -105.4744
11.3361 | -105.4836
11.3398 | -105.4952
11.3436 | -105.5068
11.3473 | -105.5222
11.351 | -105.5378
11.3547 | -105.5544
11.3584 | -105.5735
11.3622 | -105.5926
11.3659 | -105.6135
11.3696 | -105.6323
11.3733 | -105.6473
11.377 | -105.6653
11.3808 | -105.6792
11.3845 | -105.6948
11.3882 | -105.7092
11.3919 | -105.7256
11.3956 | -105.7385
11.3993 | -105.7485
11.4031 | -105.7624
11.4068 | -105.7798
11.4105 | -105.7977
11.4142 | -105.8146
11.4179 | -105.8261
11.4217 | -105.8425
11.4254 | -105.8598
11.4291 | -105.8735
11.4328 | -105.8873
11.4365 | -105.9019
11.4403 | -105.9167
11.444 | -105.93
11.4477 | -105.9442
11.4514 | -105.9583
11.4551 | -105.9722
11.4589 | -105.9892
11.4626 | -106.0089
11.4663 | -106.0254
11.47 | -106.0358
11.4737 | -106.0407
11.4775 | -106.046
11.4812 | -106.0504
11.4849 | -106.0512
11.4886 | -106.0513
11.4923 | -106.0449
11.496 | -106.035
11.4998 | -106.0281
11.5035 | -106.0184
11.5072 | -106.0087
11.5109 | -105.9981
11.5146 | -105.9857
11.5184 | -105.9743
11.5221 | -105.9634
11.5258 | -105.9494
11.5295 | -105.9316
11.5332 | -105.9172
11.537 | -105.9046
11.5407 | -105.8973
11.5444 | -105.89
11.5481 | -105.884
11.5518 | -105.8763
11.5556 | -105.8663
11.5593 | -105.8558
11.563 | -105.848
11.5667 | -105.8398
11.5704 | -105.8319
11.5742 | -105.8252
11.5779 | -105.8205
11.5816 | -105.8163
11.5853 | -105.8133
11.589 | -105.8083
11.5927 | -105.8043
11.5965 | -105.8006
11.6002 | -105.7966
11.6039 | -105.7913
11.6076 | -105.787
11.6113 | -105.7845
11.6151 | -105.78
11.6188 | -105.7753
11.6225 | -105.7713
11.6262 | -105.7678
11.6299 | -105.7663
11.6337 | -105.7626
11.6374 | -105.7627
11.6411 | -105.7644
11.6448 | -105.7611
11.6485 | -105.7581
11.6523 | -105.759
11.656 | -105.7626
11.6597 | -105.7673
11.6634 | -105.7698
11.6671 | -105.7727
11.6709 | -105.7759
11.6746 | -105.778
11.6783 | -105.7821
11.682 | -105.7843
11.6857 | -105.7866
11.6894 | -105.7894
11.6932 | -105.7979
11.6969 | -105.8077
11.7006 | -105.812
11.7043 | -105.8149
11.708 | -105.816
11.7118 | -105.8137
11.7155 | -105.8129
11.7192 | -105.8136
11.7229 | -105.8156
11.7266 | -105.8188
11.7304 | -105.8214
11.7341 | -105.8264
11.7378 | -105.8314
11.7415 | -105.8395
11.7452 | -105.8465
11.749 | -105.8552
11.7527 | -105.8717
11.7564 | -105.8921
11.7601 | -105.908
11.7638 | -105.9261
11.7675 | -105.9487
11.7713 | -105.9729
11.775 | -105.9965
11.7787 | -106.0207
11.7824 | -106.0444
11.7861 | -106.0668
11.7899 | -106.0872
11.7936 | -106.1065
11.7973 | -106.1236
11.801 | -106.1314
11.8047 | -106.1465
11.8085 | -106.1546
11.8122 | -106.1603
11.8159 | -106.1666
11.8196 | -106.1688
11.8233 | -106.1679
11.8271 | -106.1661
11.8308 | -106.1653
11.8345 | -106.1644
11.8382 | -106.1607
11.8419 | -106.1537
11.8457 | -106.1445
11.8494 | -106.1308
11.8531 | -106.1176
11.8568 | -106.1011
11.8605 | -106.0835
11.8642 | -106.0646
11.868 | -106.0456
11.8717 | -106.0265
11.8754 | -106.0057
11.8791 | -105.984
11.8828 | -105.9598
11.8866 | -105.9402
11.8903 | -105.919
11.894 | -105.8999
11.8977 | -105.8811
11.9014 | -105.8631
11.9052 | -105.8464
11.9089 | -105.8302
11.9126 | -105.8156
11.9163 | -105.808
11.92 | -105.8036
11.9238 | -105.7943
11.9275 | -105.7852
11.9312 | -105.7861
11.9349 | -105.7903
11.9386 | -105.7942
11.9424 | -105.8021
11.9461 | -105.8115
11.9498 | -105.8221
11.9535 | -105.838
11.9572 | -105.8556
11.9609 | -105.8745
11.9647 | -105.8946
11.9684 | -105.9153
11.9721 | -105.9353
11.9758 | -105.9559
11.9795 | -105.9761
11.9833 | -105.996
11.987 | -106.0141
11.9907 | -106.0274
11.9944 | -106.0412
11.9981 | -106.0487
12.0019 | -106.053
12.0056 | -106.0552
12.0093 | -106.0562
12.013 | -106.0564
12.0167 | -106.0581
12.0205 | -106.0586
12.0242 | -106.0591
12.0279 | -106.055
12.0316 | -106.0485
12.0353 | -106.0399
12.0391 | -106.0249
12.0428 | -106.0116
12.0465 | -106
12.0502 | -105.9886
12.0539 | -105.979
12.0576 | -105.9693
12.0614 | -105.9555
12.0651 | -105.9454
12.0688 | -105.9321
12.0725 | -105.9161
12.0762 | -105.9026
12.08 | -105.8929
12.0837 | -105.883
12.0874 | -105.8722
12.0911 | -105.8622
12.0948 | -105.8521
12.0986 | -105.8467
12.1023 | -105.8427
12.106 | -105.8365
12.1097 | -105.8297
12.1134 | -105.8308
12.1172 | -105.8335
12.1209 | -105.8358
12.1246 | -105.8396
12.1283 | -105.8424
12.132 | -105.8484
12.1358 | -105.8576
12.1395 | -105.8678
12.1432 | -105.8788
12.1469 | -105.8883
12.1506 | -105.8955
12.1543 | -105.9055
12.1581 | -105.9151
12.1618 | -105.9206
12.1655 | -105.9254
12.1692 | -105.9233
12.1729 | -105.9241
12.1767 | -105.9251
12.1804 | -105.922
12.1841 | -105.9198
12.1878 | -105.9173
12.1915 | -105.914
12.1953 | -105.9101
12.199 | -105.9056
12.2027 | -105.904
12.2064 | -105.9029
12.2101 | -105.9002
12.2139 | -105.901
12.2176 | -105.8977
12.2213 | -105.8914
12.225 | -105.8866
12.2287 | -105.8792
12.2325 | -105.8743
12.2362 | -105.8713
12.2399 | -105.8683
12.2436 | -105.8675
12.2473 | -105.8647
12.251 | -105.8647
12.2548 | -105.864
12.2585 | -105.8638
12.2622 | -105.8643
12.2659 | -105.8653
12.2696 | -105.8658
12.2734 | -105.8697
12.2771 | -105.87
12.2808 | -105.8733
12.2845 | -105.8784
12.2882 | -105.8784
12.292 | -105.881
12.2957 | -105.887
12.2994 | -105.8925
12.3031 | -105.8991
12.3068 | -105.9081
12.3106 | -105.9176
12.3143 | -105.9249
12.318 | -105.9341
12.3217 | -105.9432
12.3254 | -105.9535
12.3291 | -105.9628
12.3329 | -105.9671
12.3366 | -105.9746
12.3403 | -105.9858
12.344 | -105.9991
12.3477 | -106.012
12.3515 | -106.021
12.3552 | -106.027
12.3589 | -106.0332
12.3626 | -106.037
12.3663 | -106.0424
12.3701 | -106.0497
12.3738 | -106.0564
12.3775 | -106.0596
12.3812 | -106.0632
12.3849 | -106.0663
12.3887 | -106.0671
12.3924 | -106.0689
12.3961 | -106.0727
12.3998 | -106.0717
12.4035 | -106.0663
12.4073 | -106.0633
12.411 | -106.0577
12.4147 | -106.0461
12.4184 | -106.0303
12.4221 | -106.0125
12.4258 | -105.996
12.4296 | -105.9763
12.4333 | -105.9562
12.437 | -105.9401
12.4407 | -105.9228
12.4444 | -105.9029
12.4482 | -105.8847
12.4519 | -105.8666
12.4556 | -105.8523
12.4593 | -105.838
12.463 | -105.8243
12.4668 | -105.8132
12.4705 | -105.8044
12.4742 | -105.7978
12.4779 | -105.7955
12.4816 | -105.7961
12.4854 | -105.7969
12.4891 | -105.8004
12.4928 | -105.8095
12.4965 | -105.821
12.5002 | -105.8376
12.504 | -105.8555
12.5077 | -105.8768
12.5114 | -105.9019
12.5151 | -105.9264
12.5188 | -105.9461
12.5225 | -105.9634
12.5263 | -105.9804
12.53 | -106.0017
12.5337 | -106.0192
12.5374 | -106.0319
12.5411 | -106.046
12.5449 | -106.0582
12.5486 | -106.0667
12.5523 | -106.0724
12.556 | -106.0767
12.5597 | -106.0811
12.5635 | -106.0825
12.5672 | -106.0835
12.5709 | -106.0858
12.5746 | -106.0838
12.5783 | -106.0806
12.5821 | -106.078
12.5858 | -106.0744
12.5895 | -106.0678
12.5932 | -106.0655
12.5969 | -106.0666
12.6007 | -106.0696
12.6044 | -106.0773
12.6081 | -106.0824
12.6118 | -106.0869
12.6155 | -106.0922
12.6192 | -106.0997
12.623 | -106.1113
12.6267 | -106.1274
12.6304 | -106.1432
12.6341 | -106.1622
12.6378 | -106.1858
12.6416 | -106.2083
12.6453 | -106.2342
12.649 | -106.2632
12.6527 | -106.2894
12.6564 | -106.3192
12.6602 | -106.3469
12.6639 | -106.3727
12.6676 | -106.399
12.6713 | -106.422
12.675 | -106.4425
12.6788 | -106.4636
12.6825 | -106.4837
12.6862 | -106.4974
12.6899 | -106.5055
12.6936 | -106.511
12.6974 | -106.5156
12.7011 | -106.5214
12.7048 | -106.5243
12.7085 | -106.5225
12.7122 | -106.5144
12.7159 | -106.5079
12.7197 | -106.5005
12.7234 | -106.4923
12.7271 | -106.481
12.7308 | -106.4668
12.7345 | -106.4531
12.7383 | -106.4394
12.742 | -106.4279
12.7457 | -106.4164
12.7494 | -106.4024
12.7531 | -106.388
12.7569 | -106.3775
12.7606 | -106.3711
12.7643 | -106.365
12.768 | -106.3585
12.7717 | -106.3537
12.7755 | -106.3462
12.7792 | -106.3386
12.7829 | -106.3325
12.7866 | -106.3252
12.7903 | -106.3199
12.794 | -106.3147
12.7978 | -106.3077
12.8015 | -106.2975
12.8052 | -106.2867
12.8089 | -106.2768
12.8126 | -106.2721
12.8164 | -106.2618
12.8201 | -106.2472
12.8238 | -106.2358
12.8275 | -106.2254
12.8312 | -106.2114
12.835 | -106.1994
12.8387 | -106.1834
12.8424 | -106.1688
12.8461 | -106.1547
12.8498 | -106.1404
12.8536 | -106.1287
12.8573 | -106.1186
12.861 | -106.1088
12.8647 | -106.1008
12.8684 | -106.0966
12.8722 | -106.098
12.8759 | -106.1025
12.8796 | -106.1056
12.8833 | -106.1089
12.887 | -106.1087
12.8907 | -106.1123
12.8945 | -106.1182
12.8982 | -106.1267
12.9019 | -106.1345
12.9056 | -106.1438
12.9093 | -106.1464
12.9131 | -106.1497
12.9168 | -106.1552
12.9205 | -106.1577
12.9242 | -106.1607
12.9279 | -106.1608
12.9317 | -106.1577
12.9354 | -106.1572
12.9391 | -106.1577
12.9428 | -106.1541
12.9465 | -106.1496
12.9503 | -106.147
12.954 | -106.1458
12.9577 | -106.1436
12.9614 | -106.142
12.9651 | -106.1414
12.9689 | -106.1419
12.9726 | -106.1449
12.9763 | -106.1494
12.98 | -106.1527
12.9837 | -106.1561
12.9874 | -106.1622
12.9912 | -106.1689
12.9949 | -106.1723
12.9986 | -106.1728
13.0023 | -106.1771
13.006 | -106.1835
13.0098 | -106.1885
13.0135 | -106.1908
13.0172 | -106.1939
13.0209 | -106.1963
13.0246 | -106.1985
13.0284 | -106.199
13.0321 | -106.2008
13.0358 | -106.2048
13.0395 | -106.2064
13.0432 | -106.2076
13.047 | -106.2085
13.0507 | -106.2085
13.0544 | -106.2103
13.0581 | -106.2081
13.0618 | -106.2085
13.0656 | -106.2101
13.0693 | -106.209
13.073 | -106.2057
13.0767 | -106.2046
13.0804 | -106.2002
13.0841 | -106.1961
13.0879 | -106.1906
13.0916 | -106.1846
13.0953 | -106.1819
13.099 | -106.1803
13.1027 | -106.1791
13.1065 | -106.1775
13.1102 | -106.1751
13.1139 | -106.1741
13.1176 | -106.1723
13.1213 | -106.1693
13.1251 | -106.1608
13.1288 | -106.1516
13.1325 | -106.1439
13.1362 | -106.1359
13.1399 | -106.1265
13.1437 | -106.1125
13.1474 | -106.0943
13.1511 | -106.0775
13.1548 | -106.0576
13.1585 | -106.0366
13.1623 | -106.0131
13.166 | -105.9857
13.1697 | -105.9606
13.1734 | -105.9376
13.1771 | -105.9119
13.1808 | -105.8864
13.1846 | -105.8625
13.1883 | -105.844
13.192 | -105.8257
13.1957 | -105.8092
13.1994 | -105.7969
13.2032 | -105.7859
13.2069 | -105.7769
13.2106 | -105.7716
13.2143 | -105.7694
13.218 | -105.7697
13.2218 | -105.7733
13.2255 | -105.7776
13.2292 | -105.7819
13.2329 | -105.7898
13.2366 | -105.7954
13.2404 | -105.8019
13.2441 | -105.8121
13.2478 | -105.8248
13.2515 | -105.8373
13.2552 | -105.848
13.2589 | -105.8571
13.2627 | -105.8632
13.2664 | -105.8652
13.2701 | -105.8663
13.2738 | -105.8668
13.2775 | -105.8676
13.2813 | -105.8671
13.285 | -105.8599
13.2887 | -105.8495
13.2924 | -105.8374
13.2961 | -105.8218
13.2999 | -105.8055
13.3036 | -105.7873
13.3073 | -105.771
13.311 | -105.7575
13.3147 | -105.739
13.3185 | -105.7156
13.3222 | -105.6948
13.3259 | -105.6746
13.3296 | -105.6596
13.3333 | -105.6457
13.3371 | -105.6323
13.3408 | -105.6202
13.3445 | -105.6109
13.3482 | -105.6036
13.3519 | -105.5998
13.3556 | -105.6006
13.3594 | -105.6006
13.3631 | -105.6033
13.3668 | -105.609
13.3705 | -105.6161
13.3742 | -105.6256
13.378 | -105.6311
13.3817 | -105.6375
13.3854 | -105.6431
13.3891 | -105.6492
13.3928 | -105.6602
13.3966 | -105.6723
13.4003 | -105.6823
13.404 | -105.6892
13.4077 | -105.6962
13.4114 | -105.7015
13.4152 | -105.7066
13.4189 | -105.7131
13.4226 | -105.7236
13.4263 | -105.7336
13.43 | -105.7452
13.4338 | -105.7545
13.4375 | -105.7578
13.4412 | -105.7602
13.4449 | -105.7658
13.4486 | -105.7718
13.4523 | -105.7762
13.4561 | -105.786
13.4598 | -105.7965
13.4635 | -105.8041
13.4672 | -105.8069
13.4709 | -105.8072
13.4747 | -105.8094
13.4784 | -105.8139
13.4821 | -105.8205
13.4858 | -105.8277
13.4895 | -105.8341
13.4933 | -105.8405
13.497 | -105.8438
13.5007 | -105.8483
13.5044 | -105.8513
13.5081 | -105.8584
13.5119 | -105.8684
13.5156 | -105.8741
13.5193 | -105.8759
13.523 | -105.8785
13.5267 | -105.8854
13.5305 | -105.8937
13.5342 | -105.9001
13.5379 | -105.9068
13.5416 | -105.9142
13.5453 | -105.9203
13.549 | -105.9275
13.5528 | -105.9353
13.5565 | -105.9383
13.5602 | -105.9442
13.5639 | -105.9511
13.5676 | -105.9592
13.5714 | -105.969
13.5751 | -105.9806
13.5788 | -105.994
13.5825 | -106.0067
13.5862 | -106.0181
13.59 | -106.028
13.5937 | -106.0392
13.5974 | -106.0534
13.6011 | -106.0659
13.6048 | -106.0784
13.6086 | -106.0937
13.6123 | -106.1098
13.616 | -106.1254
13.6197 | -106.1387
13.6234 | -106.1519
13.6272 | -106.165
13.6309 | -106.1791
13.6346 | -106.197
13.6383 | -106.2144
13.642 | -106.2252
13.6457 | -106.2382
13.6495 | -106.2506
13.6532 | -106.2565
13.6569 | -106.2626
13.6606 | -106.2701
13.6643 | -106.2779
13.6681 | -106.2848
13.6718 | -106.29
13.6755 | -106.2934
13.6792 | -106.2967
13.6829 | -106.2985
13.6867 | -106.301
13.6904 | -106.3023
13.6941 | -106.3023
13.6978 | -106.2974
13.7015 | -106.296
13.7053 | -106.297
13.709 | -106.295
13.7127 | -106.2931
13.7164 | -106.2932
13.7201 | -106.2878
13.7238 | -106.2834
13.7276 | -106.2823
13.7313 | -106.2811
13.735 | -106.2765
13.7387 | -106.2726
13.7424 | -106.2722
13.7462 | -106.2717
13.7499 | -106.2723
13.7536 | -106.2764
13.7573 | -106.2771
13.761 | -106.2769
13.7648 | -106.2799
13.7685 | -106.2839
13.7722 | -106.2885
13.7759 | -106.2929
13.7796 | -106.2984
13.7834 | -106.3075
13.7871 | -106.317
13.7908 | -106.3255
13.7945 | -106.3358
13.7982 | -106.3479
13.802 | -106.3589
13.8057 | -106.3675
13.8094 | -106.3737
13.8131 | -106.3793
13.8168 | -106.3881
13.8205 | -106.3951
13.8243 | -106.403
13.828 | -106.4092
13.8317 | -106.4102
13.8354 | -106.4098
13.8391 | -106.4093
13.8429 | -106.405
13.8466 | -106.4023
13.8503 | -106.4007
13.854 | -106.3977
13.8577 | -106.3926
13.8615 | -106.386
13.8652 | -106.3788
13.8689 | -106.3734
13.8726 | -106.3692
13.8763 | -106.3646
13.8801 | -106.3594
13.8838 | -106.3574
13.8875 | -106.3583
13.8912 | -106.3565
13.8949 | -106.3585
13.8987 | -106.3622
13.9024 | -106.3664
13.9061 | -106.3732
13.9098 | -106.3826
13.9135 | -106.3923
13.9172 | -106.4057
13.921 | -106.4162
13.9247 | -106.4269
13.9284 | -106.4405
13.9321 | -106.4564
13.9358 | -106.4711
13.9396 | -106.4848
13.9433 | -106.5
13.947 | -106.5095
13.9507 | -106.5145
13.9544 | -106.5168
13.9582 | -106.5149
13.9619 | -106.5089
13.9656 | -106.4967
13.9693 | -106.4863
13.973 | -106.4731
13.9768 | -106.4585
13.9805 | -106.4386
13.9842 | -106.4144
13.9879 | -106.3891
13.9916 | -106.36
13.9954 | -106.3251
13.9991 | -106.2879
14.0028 | -106.2451
14.0065 | -106.201
14.0102 | -106.1582
14.0139 | -106.1131
14.0177 | -106.0666
14.0214 | -106.0216
14.0251 | -105.9784
14.0288 | -105.9374
14.0325 | -105.8979
14.0363 | -105.8588
14.04 | -105.825
14.0437 | -105.796
14.0474 | -105.7699
14.0511 | -105.7491
14.0549 | -105.7357
14.0586 | -105.7286
14.0623 | -105.7253
14.066 | -105.7258
14.0697 | -105.7345
14.0735 | -105.7482
14.0772 | -105.7668
14.0809 | -105.7939
14.0846 | -105.824
14.0883 | -105.8548
14.0921 | -105.8919
14.0958 | -105.9278
14.0995 | -105.9636
14.1032 | -106.0054
14.1069 | -106.0471
14.1106 | -106.0902
14.1144 | -106.1341
14.1181 | -106.1756
14.1218 | -106.2177
14.1255 | -106.2562
14.1292 | -106.2911
14.133 | -106.3259
14.1367 | -106.357
14.1404 | -106.3855
14.1441 | -106.4126
14.1478 | -106.4353
14.1516 | -106.4596
14.1553 | -106.4847
14.159 | -106.5088
14.1627 | -106.5341
14.1664 | -106.5574
14.1702 | -106.5761
14.1739 | -106.5896
14.1776 | -106.601
14.1813 | -106.6092
14.185 | -106.6208
14.1887 | -106.6294
14.1925 | -106.6358
14.1962 | -106.6441
14.1999 | -106.6515
14.2036 | -106.6582
14.2073 | -106.665
14.2111 | -106.6675
14.2148 | -106.6639
14.2185 | -106.6601
14.2222 | -106.6513
14.2259 | -106.6397
14.2297 | -106.6281
14.2334 | -106.6175
14.2371 | -106.6058
14.2408 | -106.5935
14.2445 | -106.5759
14.2483 | -106.5535
14.252 | -106.5315
14.2557 | -106.5071
14.2594 | -106.4794
14.2631 | -106.4526
14.2669 | -106.429
14.2706 | -106.4073
14.2743 | -106.38
14.278 | -106.3542
14.2817 | -106.3316
14.2854 | -106.3056
14.2892 | -106.2831
14.2929 | -106.2685
14.2966 | -106.2558
14.3003 | -106.2466
14.304 | -106.2382
14.3078 | -106.2295
14.3115 | -106.2239
14.3152 | -106.222
14.3189 | -106.2247
14.3226 | -106.2289
14.3264 | -106.2363
14.3301 | -106.249
14.3338 | -106.2658
14.3375 | -106.2823
14.3412 | -106.3019
14.345 | -106.3226
14.3487 | -106.3412
14.3524 | -106.36
14.3561 | -106.3795
14.3598 | -106.402
14.3636 | -106.4235
14.3673 | -106.4409
14.371 | -106.4588
14.3747 | -106.4772
14.3784 | -106.4934
14.3821 | -106.5073
14.3859 | -106.5204
14.3896 | -106.5302
14.3933 | -106.5405
14.397 | -106.5501
14.4007 | -106.5543
14.4045 | -106.5568
14.4082 | -106.5565
14.4119 | -106.5576
14.4156 | -106.5592
14.4193 | -106.5579
14.4231 | -106.5549
14.4268 | -106.5522
14.4305 | -106.5478
14.4342 | -106.5424
14.4379 | -106.5346
14.4417 | -106.5277
14.4454 | -106.5193
14.4491 | -106.5121
14.4528 | -106.5079
14.4565 | -106.5021
14.4603 | -106.4978
14.464 | -106.4945
14.4677 | -106.4915
14.4714 | -106.4897
14.4751 | -106.4886
14.4788 | -106.4881
14.4826 | -106.4887
14.4863 | -106.4943
14.49 | -106.4965
14.4937 | -106.5004
14.4974 | -106.5039
14.5012 | -106.5057
14.5049 | -106.5103
14.5086 | -106.5164
14.5123 | -106.5222
14.516 | -106.5313
14.5198 | -106.5419
14.5235 | -106.5496
14.5272 | -106.5585
14.5309 | -106.5698
14.5346 | -106.5797
14.5384 | -106.5898
14.5421 | -106.5995
14.5458 | -106.6101
14.5495 | -106.6205
14.5532 | -106.6317
14.557 | -106.6466
14.5607 | -106.6612
14.5644 | -106.6732
14.5681 | -106.6814
14.5718 | -106.6899
14.5755 | -106.7027
14.5793 | -106.7131
14.583 | -106.7232
14.5867 | -106.7307
14.5904 | -106.7411
14.5941 | -106.7513
14.5979 | -106.7601
14.6016 | -106.7693
14.6053 | -106.7768
14.609 | -106.7876
14.6127 | -106.7983
14.6165 | -106.8062
14.6202 | -106.8162
14.6239 | -106.8256
14.6276 | -106.8364
14.6313 | -106.8474
14.6351 | -106.8585
14.6388 | -106.8677
14.6425 | -106.8754
14.6462 | -106.8851
14.6499 | -106.8998
14.6536 | -106.9174
14.6574 | -106.9346
14.6611 | -106.9502
14.6648 | -106.9629
14.6685 | -106.9745
14.6722 | -106.9836
14.676 | -106.9914
14.6797 | -106.9989
14.6834 | -107.0094
14.6871 | -107.0222
14.6908 | -107.0326
14.6946 | -107.042
14.6983 | -107.049
14.702 | -107.0541
14.7057 | -107.0608
14.7094 | -107.067
14.7132 | -107.0718
14.7169 | -107.0755
14.7206 | -107.0751
14.7243 | -107.0751
14.728 | -107.0716
14.7318 | -107.0669
14.7355 | -107.062
14.7392 | -107.0556
14.7429 | -107.0462
14.7466 | -107.0347
14.7503 | -107.0236
14.7541 | -107.0129
14.7578 | -107.0001
14.7615 | -106.9872
14.7652 | -106.9719
14.7689 | -106.9536
14.7727 | -106.9373
14.7764 | -106.9232
14.7801 | -106.9058
14.7838 | -106.8895
14.7875 | -106.8704
14.7913 | -106.8536
14.795 | -106.8372
14.7987 | -106.8203
14.8024 | -106.8017
14.8061 | -106.785
14.8099 | -106.7683
14.8136 | -106.7487
14.8173 | -106.7298
14.821 | -106.7114
14.8247 | -106.6912
14.8285 | -106.676
14.8322 | -106.6563
14.8359 | -106.6352
14.8396 | -106.6161
14.8433 | -106.5956
14.847 | -106.5742
14.8508 | -106.556
14.8545 | -106.5393
14.8582 | -106.5195
14.8619 | -106.4981
14.8656 | -106.4771
14.8694 | -106.4585
14.8731 | -106.4412
14.8768 | -106.4266
14.8805 | -106.4141
14.8842 | -106.4035
14.888 | -106.3941
14.8917 | -106.3865
14.8954 | -106.3782
14.8991 | -106.3708
14.9028 | -106.368
14.9066 | -106.3641
14.9103 | -106.3617
14.914 | -106.3622
14.9177 | -106.3638
14.9214 | -106.3659
14.9252 | -106.3699
14.9289 | -106.3738
14.9326 | -106.3798
14.9363 | -106.3861
14.94 | -106.392
14.9437 | -106.3956
14.9475 | -106.3986
14.9512 | -106.4023
14.9549 | -106.4075
14.9586 | -106.411
14.9623 | -106.4121
14.9661 | -106.4135
14.9698 | -106.4149
14.9735 | -106.4151
14.9772 | -106.4173
14.9809 | -106.4165
14.9847 | -106.4126
14.9884 | -106.4067
14.9921 | -106.4043
14.9958 | -106.4014
14.9995 | -106.399
15.0033 | -106.3965
15.007 | -106.3943
15.0107 | -106.3921
15.0144 | -106.3889
15.0181 | -106.385
15.0219 | -106.3856
15.0256 | -106.3828
15.0293 | -106.3812
15.033 | -106.3816
15.0367 | -106.3812
15.0404 | -106.3814
15.0442 | -106.3801
15.0479 | -106.3798
15.0516 | -106.3826
15.0553 | -106.3871
15.059 | -106.3911
15.0628 | -106.3968
15.0665 | -106.4025
15.0702 | -106.4055
15.0739 | -106.4085
15.0776 | -106.4128
15.0814 | -106.4183
15.0851 | -106.4218
15.0888 | -106.4256
15.0925 | -106.4315
15.0962 | -106.4371
15.1 | -106.4441
15.1037 | -106.4494
15.1074 | -106.4498
15.1111 | -106.4493
15.1148 | -106.4494
15.1186 | -106.4472
15.1223 | -106.4468
15.126 | -106.4448
15.1297 | -106.4431
15.1334 | -106.442
15.1371 | -106.441
15.1409 | -106.4392
15.1446 | -106.438
15.1483 | -106.4378
15.152 | -106.439
15.1557 | -106.4393
15.1595 | -106.4411
15.1632 | -106.4422
15.1669 | -106.4402
15.1706 | -106.4379
15.1743 | -106.4355
15.1781 | -106.4321
15.1818 | -106.4305
15.1855 | -106.4348
15.1892 | -106.4373
15.1929 | -106.4371
15.1967 | -106.4412
15.2004 | -106.4487
15.2041 | -106.4513
15.2078 | -106.4529
15.2115 | -106.4573
15.2152 | -106.4621
15.219 | -106.4658
15.2227 | -106.4674
15.2264 | -106.4714
15.2301 | -106.4764
15.2338 | -106.4805
15.2376 | -106.4838
15.2413 | -106.4883
15.245 | -106.4938
15.2487 | -106.4986
15.2524 | -106.5038
15.2562 | -106.5095
15.2599 | -106.5117
15.2636 | -106.5129
15.2673 | -106.5192
15.271 | -106.5223
15.2748 | -106.5235
15.2785 | -106.5261
15.2822 | -106.5284
15.2859 | -106.5329
15.2896 | -106.538
15.2934 | -106.5412
15.2971 | -106.5446
15.3008 | -106.548
15.3045 | -106.5513
15.3082 | -106.5511
15.3119 | -106.5506
15.3157 | -106.5498
15.3194 | -106.5471
15.3231 | -106.5448
15.3268 | -106.5447
15.3305 | -106.5437
15.3343 | -106.5428
15.338 | -106.543
15.3417 | -106.5419
15.3454 | -106.541
15.3491 | -106.5379
15.3529 | -106.5322
15.3566 | -106.5272
15.3603 | -106.5223
15.364 | -106.5169
15.3677 | -106.5139
15.3715 | -106.5119
15.3752 | -106.5111
15.3789 | -106.509
15.3826 | -106.5078
15.3863 | -106.5058
15.3901 | -106.5047
15.3938 | -106.504
15.3975 | -106.5013
15.4012 | -106.4964
15.4049 | -106.4877
15.4086 | -106.479
15.4124 | -106.4744
15.4161 | -106.4694
15.4198 | -106.4629
15.4235 | -106.4597
15.4272 | -106.4552
15.431 | -106.4483
15.4347 | -106.4406
15.4384 | -106.4345
15.4421 | -106.4291
15.4458 | -106.4236
15.4496 | -106.4165
15.4533 | -106.4098
15.457 | -106.4042
15.4607 | -106.3973
15.4644 | -106.3903
15.4682 | -106.3844
15.4719 | -106.3761
15.4756 | -106.3661
15.4793 | -106.3599
15.483 | -106.3525
15.4868 | -106.3429
15.4905 | -106.3393
15.4942 | -106.3368
15.4979 | -106.3354
15.5016 | -106.3313
15.5053 | -106.3256
15.5091 | -106.3203
15.5128 | -106.3189
15.5165 | -106.3147
15.5202 | -106.3118
15.5239 | -106.3097
15.5277 | -106.3068
15.5314 | -106.3044
15.5351 | -106.3014
15.5388 | -106.3
15.5425 | -106.2991
15.5463 | -106.2963
15.55 | -106.2945
15.5537 | -106.2928
15.5574 | -106.2884
15.5611 | -106.2833
15.5649 | -106.2808
15.5686 | -106.2787
15.5723 | -106.2792
15.576 | -106.278
15.5797 | -106.2773
15.5834 | -106.2744
15.5872 | -106.2746
15.5909 | -106.2747
15.5946 | -106.2747
15.5983 | -106.2747
15.602 | -106.2755
15.6058 | -106.2774
15.6095 | -106.2787
15.6132 | -106.2771
15.6169 | -106.2771
15.6206 | -106.2807
15.6244 | -106.2829
15.6281 | -106.2861
15.6318 | -106.2909
15.6355 | -106.2924
15.6392 | -106.2935
15.643 | -106.2939
15.6467 | -106.2955
15.6504 | -106.2968
15.6541 | -106.2979
15.6578 | -106.3028
15.6616 | -106.3062
15.6653 | -106.3107
15.669 | -106.3172
15.6727 | -106.3235
15.6764 | -106.3304
15.6801 | -106.3348
15.6839 | -106.3374
15.6876 | -106.3428
15.6913 | -106.3492
15.695 | -106.3571
15.6987 | -106.3666
15.7025 | -106.375
15.7062 | -106.3822
15.7099 | -106.3834
15.7136 | -106.3816
15.7173 | -106.3792
15.7211 | -106.3764
15.7248 | -106.3758
15.7285 | -106.3737
15.7322 | -106.3703
15.7359 | -106.3649
15.7397 | -106.358
15.7434 | -106.3503
15.7471 | -106.3401
15.7508 | -106.3304
15.7545 | -106.3208
15.7583 | -106.3102
15.762 | -106.2969
15.7657 | -106.2842
15.7694 | -106.2721
15.7731 | -106.2586
15.7768 | -106.2437
15.7806 | -106.229
15.7843 | -106.2136
15.788 | -106.2006
15.7917 | -106.1878
15.7954 | -106.1768
15.7992 | -106.1645
15.8029 | -106.1548
15.8066 | -106.1461
15.8103 | -106.1407
15.814 | -106.1343
15.8178 | -106.1284
15.8215 | -106.1245
15.8252 | -106.1249
15.8289 | -106.129
15.8326 | -106.1355
15.8364 | -106.1462
15.8401 | -106.1594
15.8438 | -106.1739
15.8475 | -106.1883
15.8512 | -106.2036
15.855 | -106.2179
15.8587 | -106.2325
15.8624 | -106.2507
15.8661 | -106.2695
15.8698 | -106.288
15.8735 | -106.3045
15.8773 | -106.324
15.881 | -106.3445
15.8847 | -106.3619
15.8884 | -106.377
15.8921 | -106.3889
15.8959 | -106.3974
15.8996 | -106.4071
15.9033 | -106.4132
15.907 | -106.4185
15.9107 | -106.4219
15.9145 | -106.4206
15.9182 | -106.4167
15.9219 | -106.411
15.9256 | -106.4032
15.9293 | -106.3938
15.9331 | -106.3852
15.9368 | -106.3785
15.9405 | -106.3686
15.9442 | -106.3551
15.9479 | -106.3417
15.9517 | -106.3296
15.9554 | -106.3169
15.9591 | -106.3063
15.9628 | -106.2926
15.9665 | -106.2786
15.9702 | -106.2634
15.974 | -106.2528
15.9777 | -106.2455
15.9814 | -106.2361
15.9851 | -106.2269
15.9888 | -106.2215
15.9926 | -106.2168
15.9963 | -106.2152
16 | -106.2137
16.0037 | -106.2109
16.0074 | -106.2121
16.0112 | -106.2138
16.0149 | -106.2122
16.0186 | -106.212
16.0223 | -106.2099
16.026 | -106.2064
16.0298 | -106.2055
16.0335 | -106.2046
16.0372 | -106.2058
16.0409 | -106.2062
16.0446 | -106.2046
16.0483 | -106.2029
16.0521 | -106.1979
16.0558 | -106.1907
16.0595 | -106.1831
16.0632 | -106.1782
16.0669 | -106.1725
16.0707 | -106.1634
16.0744 | -106.1547
16.0781 | -106.1504
16.0818 | -106.1489
16.0855 | -106.1461
16.0893 | -106.1409
16.093 | -106.1357
16.0967 | -106.1332
16.1004 | -106.1315
16.1041 | -106.1306
16.1079 | -106.1291
16.1116 | -106.1264
16.1153 | -106.126
16.119 | -106.1302
16.1227 | -106.1354
16.1265 | -106.1412
16.1302 | -106.1471
16.1339 | -106.1544
16.1376 | -106.1626
16.1413 | -106.1728
16.145 | -106.1848
16.1488 | -106.196
16.1525 | -106.2083
16.1562 | -106.2237
16.1599 | -106.2396
16.1636 | -106.2565
16.1674 | -106.2757
16.1711 | -106.2963
16.1748 | -106.3188
16.1785 | -106.3407
16.1822 | -106.3629
16.186 | -106.3865
16.1897 | -106.4071
16.1934 | -106.4303
16.1971 | -106.4536
16.2008 | -106.4763
16.2046 | -106.4978
16.2083 | -106.5189
16.212 | -106.5382
16.2157 | -106.5576
16.2194 | -106.5737
16.2232 | -106.5906
16.2269 | -106.6062
16.2306 | -106.6221
16.2343 | -106.6397
16.238 | -106.6529
16.2417 | -106.6635
16.2455 | -106.6765
16.2492 | -106.687
16.2529 | -106.6985
16.2566 | -106.7105
16.2603 | -106.7188
16.2641 | -106.7261
16.2678 | -106.7303
16.2715 | -106.7332
16.2752 | -106.7388
16.2789 | -106.7419
16.2827 | -106.7435
16.2864 | -106.7473
16.2901 | -106.7481
16.2938 | -106.7417
16.2975 | -106.7359
16.3013 | -106.7317
16.305 | -106.7268
16.3087 | -106.7194
16.3124 | -106.7101
16.3161 | -106.7002
16.3199 | -106.6903
16.3236 | -106.6821
16.3273 | -106.6747
16.331 | -106.6653
16.3347 | -106.6543
16.3384 | -106.6421
16.3422 | -106.6297
16.3459 | -106.6195
16.3496 | -106.6097
16.3533 | -106.5999
16.357 | -106.5911
16.3608 | -106.5851
16.3645 | -106.5806
16.3682 | -106.5791
16.3719 | -106.5782
16.3756 | -106.5781
16.3794 | -106.5769
16.3831 | -106.5784
16.3868 | -106.583
16.3905 | -106.5851
16.3942 | -106.5901
16.398 | -106.5956
16.4017 | -106.5985
16.4054 | -106.6029
16.4091 | -106.6102
16.4128 | -106.6202
16.4166 | -106.6275
16.4203 | -106.6329
16.424 | -106.6419
16.4277 | -106.6527
16.4314 | -106.6625
16.4351 | -106.6685
16.4389 | -106.6734
16.4426 | -106.6779
16.4463 | -106.6847
16.45 | -106.6924
16.4537 | -106.6997
16.4575 | -106.7023
16.4612 | -106.7045
16.4649 | -106.7052
16.4686 | -106.7045
16.4723 | -106.7042
16.4761 | -106.7033
16.4798 | -106.7019
16.4835 | -106.6995
16.4872 | -106.6967
16.4909 | -106.695
16.4947 | -106.6939
16.4984 | -106.6924
16.5021 | -106.6935
16.5058 | -106.6941
16.5095 | -106.6919
16.5132 | -106.6935
16.517 | -106.696
16.5207 | -106.6973
16.5244 | -106.7011
16.5281 | -106.7085
16.5318 | -106.7172
16.5356 | -106.7269
16.5393 | -106.7393
16.543 | -106.7502
16.5467 | -106.7624
16.5504 | -106.7781
16.5542 | -106.7948
16.5579 | -106.8096
16.5616 | -106.8258
16.5653 | -106.8433
16.569 | -106.8627
16.5728 | -106.8833
16.5765 | -106.9036
16.5802 | -106.9251
16.5839 | -106.9472
16.5876 | -106.9695
16.5914 | -106.9909
16.5951 | -107.0082
16.5988 | -107.0266
16.6025 | -107.0439
16.6062 | -107.0628
16.6099 | -107.08
16.6137 | -107.0965
16.6174 | -107.1121
16.6211 | -107.124
16.6248 | -107.1385
16.6285 | -107.1518
16.6323 | -107.1583
16.636 | -107.1671
16.6397 | -107.1769
16.6434 | -107.1858
16.6471 | -107.1887
16.6509 | -107.1892
16.6546 | -107.1956
16.6583 | -107.2024
16.662 | -107.2077
16.6657 | -107.2077
16.6695 | -107.2043
16.6732 | -107.2031
16.6769 | -107.2015
16.6806 | -107.1975
16.6843 | -107.1935
16.6881 | -107.1883
16.6918 | -107.1834
16.6955 | -107.18
16.6992 | -107.1736
16.7029 | -107.1655
16.7066 | -107.1558
16.7104 | -107.1475
16.7141 | -107.1384
16.7178 | -107.1247
16.7215 | -107.1098
16.7252 | -107.0949
16.729 | -107.0797
16.7327 | -107.0667
16.7364 | -107.0527
16.7401 | -107.0346
16.7438 | -107.0191
16.7476 | -107.0056
16.7513 | -106.9881
16.755 | -106.9711
16.7587 | -106.9549
16.7624 | -106.9376
16.7662 | -106.9197
16.7699 | -106.9025
16.7736 | -106.8857
16.7773 | -106.8671
16.781 | -106.8499
16.7848 | -106.8342
16.7885 | -106.8182
16.7922 | -106.8034
16.7959 | -106.7906
16.7996 | -106.7788
16.8033 | -106.7676
16.8071 | -106.7568
16.8108 | -106.7497
16.8145 | -106.743
16.8182 | -106.7352
16.8219 | -106.7274
16.8257 | -106.7225
16.8294 | -106.7213
16.8331 | -106.7229
16.8368 | -106.7229
16.8405 | -106.7215
16.8443 | -106.7207
16.848 | -106.7229
16.8517 | -106.7262
16.8554 | -106.7311
16.8591 | -106.737
16.8629 | -106.7419
16.8666 | -106.7481
16.8703 | -106.7544
16.874 | -106.7612
16.8777 | -106.7699
16.8815 | -106.7802
16.8852 | -106.788
16.8889 | -106.7972
16.8926 | -106.8074
16.8963 | -106.8124
16.9 | -106.8189
16.9038 | -106.8273
16.9075 | -106.8381
16.9112 | -106.8518
16.9149 | -106.8635
16.9186 | -106.8725
16.9224 | -106.8829
16.9261 | -106.8906
16.9298 | -106.8997
16.9335 | -106.9072
16.9372 | -106.9129
16.941 | -106.9172
16.9447 | -106.9247
16.9484 | -106.9341
16.9521 | -106.9407
16.9558 | -106.9444
16.9596 | -106.9487
16.9633 | -106.9539
16.967 | -106.9588
16.9707 | -106.9616
16.9744 | -106.9616
16.9781 | -106.9634
16.9819 | -106.9675
16.9856 | -106.9697
16.9893 | -106.9704
16.993 | -106.9729
16.9967 | -106.9731
17.0005 | -106.9729
17.0042 | -106.9723
17.0079 | -106.9706
17.0116 | -106.9677
17.0153 | -106.9627
17.0191 | -106.9587
17.0228 | -106.9568
17.0265 | -106.9552
17.0302 | -106.9524
17.0339 | -106.9461
17.0377 | -106.9401
17.0414 | -106.9358
17.0451 | -106.9314
17.0488 | -106.9254
17.0525 | -106.918
17.0563 | -106.9138
17.06 | -106.9102
17.0637 | -106.9071
17.0674 | -106.9016
17.0711 | -106.8976
17.0748 | -106.8946
17.0786 | -106.8927
17.0823 | -106.8943
17.086 | -106.8952
17.0897 | -106.8955
17.0934 | -106.8966
17.0972 | -106.8974
17.1009 | -106.8978
17.1046 | -106.8972
17.1083 | -106.8952
17.112 | -106.8949
17.1158 | -106.896
17.1195 | -106.8979
17.1232 | -106.8968
17.1269 | -106.898
17.1306 | -106.9014
17.1344 | -106.9034
17.1381 | -106.9039
17.1418 | -106.9035
17.1455 | -106.9007
17.1492 | -106.8979
17.153 | -106.8958
17.1567 | -106.8949
17.1604 | -106.8941
17.1641 | -106.8895
17.1678 | -106.8863
17.1715 | -106.8872
17.1753 | -106.8854
17.179 | -106.8837
17.1827 | -106.8811
17.1864 | -106.8786
17.1901 | -106.8738
17.1939 | -106.8671
17.1976 | -106.8568
17.2013 | -106.845
17.205 | -106.838
17.2087 | -106.83
17.2125 | -106.8207
17.2162 | -106.8096
17.2199 | -106.7983
17.2236 | -106.7843
17.2273 | -106.7701
17.2311 | -106.7577
17.2348 | -106.7465
17.2385 | -106.7366
17.2422 | -106.726
17.2459 | -106.7132
17.2497 | -106.7022
17.2534 | -106.6941
17.2571 | -106.6863
17.2608 | -106.6786
17.2645 | -106.6698
17.2682 | -106.6629
17.272 | -106.6563
17.2757 | -106.6499
17.2794 | -106.6484
17.2831 | -106.6473
17.2868 | -106.6433
17.2906 | -106.6421
17.2943 | -106.644
17.298 | -106.6443
17.3017 | -106.6431
17.3054 | -106.6453
17.3092 | -106.6495
17.3129 | -106.6534
17.3166 | -106.656
17.3203 | -106.6608
17.324 | -106.6694
17.3278 | -106.6783
17.3315 | -106.6868
17.3352 | -106.6934
17.3389 | -106.7003
17.3426 | -106.7088
17.3464 | -106.7187
17.3501 | -106.7302
17.3538 | -106.7411
17.3575 | -106.752
17.3612 | -106.7612
17.3649 | -106.7699
17.3687 | -106.7816
17.3724 | -106.7906
17.3761 | -106.7986
17.3798 | -106.8066
17.3835 | -106.8158
17.3873 | -106.8225
17.391 | -106.8288
17.3947 | -106.8331
17.3984 | -106.8345
17.4021 | -106.837
17.4059 | -106.8377
17.4096 | -106.8378
17.4133 | -106.8358
17.417 | -106.829
17.4207 | -106.8194
17.4245 | -106.8118
17.4282 | -106.8056
17.4319 | -106.7973
17.4356 | -106.788
17.4393 | -106.7737
17.443 | -106.7581
17.4468 | -106.7469
17.4505 | -106.7351
17.4542 | -106.7158
17.4579 | -106.6944
17.4616 | -106.6726
17.4654 | -106.6522
17.4691 | -106.6316
17.4728 | -106.6144
17.4765 | -106.5967
17.4802 | -106.5777
17.484 | -106.5572
17.4877 | -106.5356
17.4914 | -106.5174
17.4951 | -106.4997
17.4988 | -106.481
17.5026 | -106.4625
17.5063 | -106.4445
17.51 | -106.4281
17.5137 | -106.4139
17.5174 | -106.3994
17.5212 | -106.3871
17.5249 | -106.3777
17.5286 | -106.368
17.5323 | -106.3598
17.536 | -106.356
17.5397 | -106.3531
17.5435 | -106.3508
17.5472 | -106.349
17.5509 | -106.3477
17.5546 | -106.3496
17.5583 | -106.3516
17.5621 | -106.3577
17.5658 | -106.3657
17.5695 | -106.3715
17.5732 | -106.3787
17.5769 | -106.3862
17.5807 | -106.3968
17.5844 | -106.4106
17.5881 | -106.4247
17.5918 | -106.4369
17.5955 | -106.4505
17.5993 | -106.4651
17.603 | -106.4818
17.6067 | -106.496
17.6104 | -106.5101
17.6141 | -106.5249
17.6179 | -106.5363
17.6216 | -106.549
17.6253 | -106.5638
17.629 | -106.5768
17.6327 | -106.5891
17.6364 | -106.6036
17.6402 | -106.6164
17.6439 | -106.6291
17.6476 | -106.6421
17.6513 | -106.6523
17.655 | -106.6623
17.6588 | -106.6708
17.6625 | -106.6781
17.6662 | -106.6845
17.6699 | -106.6902
17.6736 | -106.6947
17.6774 | -106.697
17.6811 | -106.6996
17.6848 | -106.6984
17.6885 | -106.6989
17.6922 | -106.6989
17.696 | -106.6953
17.6997 | -106.6918
17.7034 | -106.6848
17.7071 | -106.6802
17.7108 | -106.6768
17.7146 | -106.6706
17.7183 | -106.6607
17.722 | -106.6495
17.7257 | -106.6363
17.7294 | -106.6244
17.7331 | -106.6123
17.7369 | -106.6025
17.7406 | -106.5928
17.7443 | -106.5816
17.748 | -106.5698
17.7517 | -106.558
17.7555 | -106.5442
17.7592 | -106.5324
17.7629 | -106.5222
17.7666 | -106.5097
17.7703 | -106.4972
17.7741 | -106.4863
17.7778 | -106.4743
17.7815 | -106.4615
17.7852 | -106.45
17.7889 | -106.4385
17.7927 | -106.4251
17.7964 | -106.4165
17.8001 | -106.4121
17.8038 | -106.4076
17.8075 | -106.4013
17.8113 | -106.3962
17.815 | -106.392
17.8187 | -106.3892
17.8224 | -106.3883
17.8261 | -106.3873
17.8298 | -106.3883
17.8336 | -106.3896
17.8373 | -106.3933
17.841 | -106.395
17.8447 | -106.3973
17.8484 | -106.4012
17.8522 | -106.4062
17.8559 | -106.4095
17.8596 | -106.4142
17.8633 | -106.4191
17.867 | -106.4221
17.8708 | -106.4269
17.8745 | -106.4328
17.8782 | -106.4348
17.8819 | -106.438
17.8856 | -106.4407
17.8894 | -106.4443
17.8931 | -106.449
17.8968 | -106.4519
17.9005 | -106.4527
17.9042 | -106.4533
17.9079 | -106.4523
17.9117 | -106.453
17.9154 | -106.4516
17.9191 | -106.4502
17.9228 | -106.4472
17.9265 | -106.4424
17.9303 | -106.4372
17.934 | -106.4335
17.9377 | -106.4328
17.9414 | -106.4284
17.9451 | -106.4216
17.9489 | -106.4178
17.9526 | -106.4142
17.9563 | -106.4074
17.96 | -106.4024
17.9637 | -106.3993
17.9675 | -106.3965
17.9712 | -106.396
17.9749 | -106.3984
17.9786 | -106.4011
17.9823 | -106.4021
17.9861 | -106.4021
17.9898 | -106.4046
17.9935 | -106.4052
17.9972 | -106.4079
18.0009 | -106.4117
18.0046 | -106.4171
18.0084 | -106.424
18.0121 | -106.4298
18.0158 | -106.4342
18.0195 | -106.4396
18.0232 | -106.4451
18.027 | -106.4517
18.0307 | -106.4571
18.0344 | -106.4619
18.0381 | -106.4664
18.0418 | -106.4731
18.0456 | -106.4816
18.0493 | -106.4911
18.053 | -106.4955
18.0567 | -106.4977
18.0604 | -106.5031
18.0642 | -106.5077
18.0679 | -106.5123
18.0716 | -106.5186
18.0753 | -106.5239
18.079 | -106.5281
18.0828 | -106.5336
18.0865 | -106.5401
18.0902 | -106.5447
18.0939 | -106.5479
18.0976 | -106.5541
18.1013 | -106.5601
18.1051 | -106.5652
18.1088 | -106.568
18.1125 | -106.5706
18.1162 | -106.5722
18.1199 | -106.5734
18.1237 | -106.5774
18.1274 | -106.5801
18.1311 | -106.5808
18.1348 | -106.5794
18.1385 | -106.5778
18.1423 | -106.5769
18.146 | -106.5735
18.1497 | -106.5671
18.1534 | -106.5614
18.1571 | -106.5577
18.1609 | -106.5545
18.1646 | -106.5547
18.1683 | -106.5526
18.172 | -106.551
18.1757 | -106.5513
18.1795 | -106.5543
18.1832 | -106.5576
18.1869 | -106.5595
18.1906 | -106.5607
18.1943 | -106.5613
18.198 | -106.5651
18.2018 | -106.5695
18.2055 | -106.5727
18.2092 | -106.5758
18.2129 | -106.5815
18.2166 | -106.5866
18.2204 | -106.5903
18.2241 | -106.5949
18.2278 | -106.5998
18.2315 | -106.6041
18.2352 | -106.6087
18.239 | -106.6129
18.2427 | -106.6168
18.2464 | -106.6225
18.2501 | -106.6255
18.2538 | -106.6283
18.2576 | -106.6269
18.2613 | -106.6244
18.265 | -106.6224
18.2687 | -106.6175
18.2724 | -106.6139
18.2762 | -106.6107
18.2799 | -106.6066
18.2836 | -106.6045
18.2873 | -106.5991
18.291 | -106.5922
18.2947 | -106.5864
18.2985 | -106.5798
18.3022 | -106.5692
18.3059 | -106.5568
18.3096 | -106.5432
18.3133 | -106.5321
18.3171 | -106.5231
18.3208 | -106.5109
18.3245 | -106.4987
18.3282 | -106.4866
18.3319 | -106.4719
18.3357 | -106.4592
18.3394 | -106.449
18.3431 | -106.4369
18.3468 | -106.4265
18.3505 | -106.4177
18.3543 | -106.4092
18.358 | -106.3999
18.3617 | -106.389
18.3654 | -106.3776
18.3691 | -106.3705
18.3728 | -106.3667
18.3766 | -106.3605
18.3803 | -106.3548
18.384 | -106.3487
18.3877 | -106.3431
18.3914 | -106.3373
18.3952 | -106.3346
18.3989 | -106.3337
18.4026 | -106.3333
18.4063 | -106.332
18.41 | -106.3323
18.4138 | -106.3348
18.4175 | -106.3362
18.4212 | -106.338
18.4249 | -106.3385
18.4286 | -106.3402
18.4324 | -106.3428
18.4361 | -106.3454
18.4398 | -106.3475
18.4435 | -106.3488
18.4472 | -106.3498
18.451 | -106.3502
18.4547 | -106.3514
18.4584 | -106.3528
18.4621 | -106.3546
18.4658 | -106.3552
18.4695 | -106.3563
18.4733 | -106.3565
18.477 | -106.3564
18.4807 | -106.3555
18.4844 | -106.3518
18.4881 | -106.3484
18.4919 | -106.3442
18.4956 | -106.3398
18.4993 | -106.3354
18.503 | -106.327
18.5067 | -106.3194
18.5105 | -106.312
18.5142 | -106.3052
18.5179 | -106.2984
18.5216 | -106.2926
18.5253 | -106.289
18.5291 | -106.2838
18.5328 | -106.2774
18.5365 | -106.2705
18.5402 | -106.2595
18.5439 | -106.2461
18.5477 | -106.2319
18.5514 | -106.222
18.5551 | -106.2123
18.5588 | -106.1999
18.5625 | -106.1877
18.5662 | -106.1759
18.57 | -106.1633
18.5737 | -106.1513
18.5774 | -106.1379
18.5811 | -106.1216
18.5848 | -106.1052
18.5886 | -106.0878
18.5923 | -106.068
18.596 | -106.0504
18.5997 | -106.0311
18.6034 | -106.0076
18.6072 | -105.9839
18.6109 | -105.9594
18.6146 | -105.9359
18.6183 | -105.9131
18.622 | -105.8907
18.6258 | -105.8698
18.6295 | -105.8457
18.6332 | -105.8208
18.6369 | -105.7943
18.6406 | -105.7709
18.6444 | -105.7461
18.6481 | -105.7219
18.6518 | -105.7012
18.6555 | -105.6794
18.6592 | -105.6564
18.6629 | -105.6343
18.6667 | -105.614
18.6704 | -105.5958
18.6741 | -105.577
18.6778 | -105.558
18.6815 | -105.5404
18.6853 | -105.5235
18.689 | -105.508
18.6927 | -105.4934
18.6964 | -105.4798
18.7001 | -105.4677
18.7039 | -105.4567
18.7076 | -105.446
18.7113 | -105.4355
18.715 | -105.4278
18.7187 | -105.4201
18.7225 | -105.4107
18.7262 | -105.4036
18.7299 | -105.397
18.7336 | -105.3897
18.7373 | -105.3831
18.7411 | -105.3756
18.7448 | -105.3713
18.7485 | -105.3687
18.7522 | -105.3645
18.7559 | -105.3611
18.7596 | -105.3552
18.7634 | -105.3482
18.7671 | -105.3436
18.7708 | -105.3367
18.7745 | -105.3305
18.7782 | -105.3233
18.782 | -105.3169
18.7857 | -105.3083
18.7894 | -105.298
18.7931 | -105.2853
18.7968 | -105.2758
18.8006 | -105.2675
18.8043 | -105.2594
18.808 | -105.2517
18.8117 | -105.2437
18.8154 | -105.2342
18.8192 | -105.2228
18.8229 | -105.212
18.8266 | -105.2016
18.8303 | -105.1895
18.834 | -105.177
18.8377 | -105.161
18.8415 | -105.1449
18.8452 | -105.1279
18.8489 | -105.1128
18.8526 | -105.098
18.8563 | -105.0831
18.8601 | -105.0663
18.8638 | -105.0496
18.8675 | -105.0341
18.8712 | -105.0219
18.8749 | -105.0079
18.8787 | -104.9936
18.8824 | -104.9772
18.8861 | -104.9583
18.8898 | -104.9413
18.8935 | -104.9235
18.8973 | -104.9069
18.901 | -104.8903
18.9047 | -104.8766
18.9084 | -104.8621
18.9121 | -104.8493
18.9159 | -104.8367
18.9196 | -104.8228
18.9233 | -104.8107
18.927 | -104.8004
18.9307 | -104.7895
18.9344 | -104.7797
18.9382 | -104.7696
18.9419 | -104.7619
18.9456 | -104.7561
18.9493 | -104.7483
18.953 | -104.7406
18.9568 | -104.7344
18.9605 | -104.7305
18.9642 | -104.7257
18.9679 | -104.7233
18.9716 | -104.7217
18.9754 | -104.7232
18.9791 | -104.7256
18.9828 | -104.7277
18.9865 | -104.7299
18.9902 | -104.7326
18.994 | -104.734
18.9977 | -104.7387
19.0014 | -104.7447
19.0051 | -104.749
19.0088 | -104.7522
19.0126 | -104.7573
19.0163 | -104.7628
19.02 | -104.7704
19.0237 | -104.7797
19.0274 | -104.7886
19.0311 | -104.7988
19.0349 | -104.8084
19.0386 | -104.8173
19.0423 | -104.8264
19.046 | -104.8346
19.0497 | -104.8409
19.0535 | -104.8488
19.0572 | -104.8583
19.0609 | -104.8686
19.0646 | -104.8761
19.0683 | -104.8798
19.0721 | -104.887
19.0758 | -104.8929
19.0795 | -104.8993
19.0832 | -104.9057
19.0869 | -104.9121
19.0907 | -104.9161
19.0944 | -104.9198
19.0981 | -104.925
19.1018 | -104.9318
19.1055 | -104.9375
19.1093 | -104.9441
19.113 | -104.9522
19.1167 | -104.961
19.1204 | -104.9661
19.1241 | -104.9701
19.1278 | -104.9767
19.1316 | -104.9859
19.1353 | -104.9916
19.139 | -104.996
19.1427 | -104.9996
19.1464 | -105.007
19.1502 | -105.0146
19.1539 | -105.025
19.1576 | -105.0362
19.1613 | -105.0451
19.165 | -105.0551
19.1688 | -105.0663
19.1725 | -105.0747
19.1762 | -105.0809
19.1799 | -105.0879
19.1836 | -105.095
19.1874 | -105.1013
19.1911 | -105.1074
19.1948 | -105.1148
19.1985 | -105.1217
19.2022 | -105.1286
19.206 | -105.1352
19.2097 | -105.1415
19.2134 | -105.1477
19.2171 | -105.1551
19.2208 | -105.167
19.2245 | -105.1777
19.2283 | -105.1862
19.232 | -105.1947
19.2357 | -105.204
19.2394 | -105.2123
19.2431 | -105.2222
19.2469 | -105.2315
19.2506 | -105.2416
19.2543 | -105.2537
19.258 | -105.2666
19.2617 | -105.281
19.2655 | -105.2956
19.2692 | -105.3084
19.2729 | -105.3242
19.2766 | -105.3441
19.2803 | -105.3648
19.2841 | -105.3832
19.2878 | -105.4011
19.2915 | -105.4169
19.2952 | -105.4339
19.2989 | -105.4545
19.3026 | -105.4749
19.3064 | -105.4946
19.3101 | -105.5172
19.3138 | -105.5409
19.3175 | -105.5593
19.3212 | -105.5804
19.325 | -105.6022
19.3287 | -105.6217
19.3324 | -105.6415
19.3361 | -105.6593
19.3398 | -105.6798
19.3436 | -105.699
19.3473 | -105.7134
19.351 | -105.7283
19.3547 | -105.7442
19.3584 | -105.7601
19.3622 | -105.7722
19.3659 | -105.7841
19.3696 | -105.7947
19.3733 | -105.8027
19.377 | -105.8075
19.3808 | -105.8111
19.3845 | -105.8156
19.3882 | -105.817
19.3919 | -105.8169
19.3956 | -105.8188
19.3993 | -105.8192
19.4031 | -105.8136
19.4068 | -105.8066
19.4105 | -105.7989
19.4142 | -105.7919
19.4179 | -105.7852
19.4217 | -105.7764
19.4254 | -105.7661
19.4291 | -105.7516
19.4328 | -105.7393
19.4365 | -105.7283
19.4403 | -105.7161
19.444 | -105.7034
19.4477 | -105.6931
19.4514 | -105.6824
19.4551 | -105.67
19.4589 | -105.6581
19.4626 | -105.6464
19.4663 | -105.6356
19.47 | -105.6252
19.4737 | -105.6148
19.4775 | -105.6052
19.4812 | -105.5965
19.4849 | -105.5904
19.4886 | -105.5841
19.4923 | -105.576
19.496 | -105.5698
19.4998 | -105.5662
19.5035 | -105.5661
19.5072 | -105.5654
19.5109 | -105.5667
19.5146 | -105.5677
19.5184 | -105.5688
19.5221 | -105.5719
19.5258 | -105.5767
19.5295 | -105.5807
19.5332 | -105.5859
19.537 | -105.5919
19.5407 | -105.6008
19.5444 | -105.6086
19.5481 | -105.615
19.5518 | -105.6213
19.5556 | -105.6289
19.5593 | -105.637
19.563 | -105.646
19.5667 | -105.6561
19.5704 | -105.6625
19.5742 | -105.6688
19.5779 | -105.6772
19.5816 | -105.6872
19.5853 | -105.6954
19.589 | -105.7034
19.5927 | -105.7115
19.5965 | -105.7223
19.6002 | -105.7315
19.6039 | -105.739
19.6076 | -105.7479
19.6113 | -105.7573
19.6151 | -105.7673
19.6188 | -105.7776
19.6225 | -105.7872
19.6262 | -105.7961
19.6299 | -105.8058
19.6337 | -105.8153
19.6374 | -105.8253
19.6411 | -105.8377
19.6448 | -105.8511
19.6485 | -105.8625
19.6523 | -105.8744
19.656 | -105.8872
19.6597 | -105.8985
19.6634 | -105.9058
19.6671 | -105.9145
19.6709 | -105.9221
19.6746 | -105.9281
19.6783 | -105.9339
19.682 | -105.9413
19.6857 | -105.9475
19.6894 | -105.9544
19.6932 | -105.9586
19.6969 | -105.9611
19.7006 | -105.9627
19.7043 | -105.9626
19.708 | -105.9638
19.7118 | -105.9643
19.7155 | -105.9622
19.7192 | -105.9616
19.7229 | -105.9585
19.7266 | -105.9542
19.7304 | -105.9501
19.7341 | -105.9486
19.7378 | -105.9439
19.7415 | -105.9381
19.7452 | -105.9303
19.749 | -105.9216
19.7527 | -105.9153
19.7564 | -105.9068
19.7601 | -105.8958
19.7638 | -105.885
19.7675 | -105.8756
19.7713 | -105.8655
19.775 | -105.8548
19.7787 | -105.846
19.7824 | -105.8337
19.7861 | -105.8211
19.7899 | -105.8136
19.7936 | -105.807
19.7973 | -105.8012
19.801 | -105.7952
19.8047 | -105.789
19.8085 | -105.784
19.8122 | -105.7796
19.8159 | -105.776
19.8196 | -105.7707
19.8233 | -105.7682
19.8271 | -105.7676
19.8308 | -105.7675
19.8345 | -105.7668
19.8382 | -105.7654
19.8419 | -105.7631
19.8457 | -105.764
19.8494 | -105.7681
19.8531 | -105.7717
19.8568 | -105.7749
19.8605 | -105.7765
19.8642 | -105.7783
19.868 | -105.7795
19.8717 | -105.7806
19.8754 | -105.7839
19.8791 | -105.786
19.8828 | -105.7876
19.8866 | -105.7902
19.8903 | -105.7931
19.894 | -105.7956
19.8977 | -105.7985
19.9014 | -105.7974
19.9052 | -105.7962
19.9089 | -105.797
19.9126 | -105.798
19.9163 | -105.7985
19.92 | -105.8003
19.9238 | -105.801
19.9275 | -105.8005
19.9312 | -105.8009
19.9349 | -105.8012
19.9386 | -105.8016
19.9424 | -105.8017
19.9461 | -105.8032
19.9498 | -105.8018
19.9535 | -105.8022
19.9572 | -105.8047
19.9609 | -105.8061
19.9647 | -105.8061
19.9684 | -105.8055
19.9721 | -105.8036
19.9758 | -105.8018
19.9795 | -105.7996
19.9833 | -105.798
19.987 | -105.8
19.9907 | -105.8027
19.9944 | -105.8021
19.9981 | -105.803
20.0019 | -105.8025
20.0056 | -105.8045
20.0093 | -105.8072
20.013 | -105.8079
20.0167 | -105.8075
20.0205 | -105.809
20.0242 | -105.8109
20.0279 | -105.8146
20.0316 | -105.818
20.0353 | -105.8198
20.0391 | -105.8227
20.0428 | -105.8277
20.0465 | -105.8307
20.0502 | -105.8331
20.0539 | -105.8381
20.0576 | -105.8413
20.0614 | -105.8426
20.0651 | -105.848
20.0688 | -105.8534
20.0725 | -105.8537
20.0762 | -105.8553
20.08 | -105.8575
20.0837 | -105.8595
20.0874 | -105.8635
20.0911 | -105.8652
20.0948 | -105.8662
20.0986 | -105.8681
20.1023 | -105.8709
20.106 | -105.872
20.1097 | -105.872
20.1134 | -105.872
20.1172 | -105.8731
20.1209 | -105.874
20.1246 | -105.8769
20.1283 | -105.8773
20.132 | -105.8748
20.1358 | -105.8696
20.1395 | -105.8618
20.1432 | -105.8565
20.1469 | -105.8514
20.1506 | -105.8431
20.1543 | -105.8354
20.1581 | -105.8274
20.1618 | -105.8187
20.1655 | -105.8091
20.1692 | -105.7991
20.1729 | -105.788
20.1767 | -105.7754
20.1804 | -105.7616
20.1841 | -105.7444
20.1878 | -105.7311
20.1915 | -105.7182
20.1953 | -105.7028
20.199 | -105.6862
20.2027 | -105.6693
20.2064 | -105.6501
20.2101 | -105.6298
20.2139 | -105.6116
20.2176 | -105.5941
20.2213 | -105.5755
20.225 | -105.5574
20.2287 | -105.5401
20.2325 | -105.5234
20.2362 | -105.5061
20.2399 | -105.4864
20.2436 | -105.465
20.2473 | -105.4461
20.251 | -105.4272
20.2548 | -105.4081
20.2585 | -105.3902
20.2622 | -105.3722
20.2659 | -105.3528
20.2696 | -105.3322
20.2734 | -105.3121
20.2771 | -105.292
20.2808 | -105.2722
20.2845 | -105.2529
20.2882 | -105.2327
20.292 | -105.211
20.2957 | -105.1921
20.2994 | -105.1758
20.3031 | -105.1596
20.3068 | -105.1441
20.3106 | -105.1295
20.3143 | -105.1174
20.318 | -105.1057
20.3217 | -105.0942
20.3254 | -105.0818
20.3291 | -105.0696
20.3329 | -105.0593
20.3366 | -105.0501
20.3403 | -105.0414
20.344 | -105.0347
20.3477 | -105.0315
20.3515 | -105.0279
20.3552 | -105.0261
20.3589 | -105.0257
20.3626 | -105.026
20.3663 | -105.0251
20.3701 | -105.0281
20.3738 | -105.0296
20.3775 | -105.0323
20.3812 | -105.0354
20.3849 | -105.0365
20.3887 | -105.0432
20.3924 | -105.0471
20.3961 | -105.0518
20.3998 | -105.0582
20.4035 | -105.0651
20.4073 | -105.0725
20.411 | -105.0791
20.4147 | -105.0844
20.4184 | -105.0891
20.4221 | -105.096
20.4258 | -105.106
20.4296 | -105.1152
20.4333 | -105.1206
20.437 | -105.1258
20.4407 | -105.133
20.4444 | -105.1386
20.4482 | -105.1455
20.4519 | -105.1498
20.4556 | -105.1543
20.4593 | -105.1596
20.463 | -105.1647
20.4668 | -105.1692
20.4705 | -105.1723
20.4742 | -105.176
20.4779 | -105.1774
20.4816 | -105.1775
20.4854 | -105.1808
20.4891 | -105.1812
20.4928 | -105.1808
20.4965 | -105.1827
20.5002 | -105.1862
20.504 | -105.186
20.5077 | -105.1839
20.5114 | -105.1803
20.5151 | -105.1771
20.5188 | -105.1735
20.5225 | -105.1681
20.5263 | -105.1611
20.53 | -105.1545
20.5337 | -105.1505
20.5374 | -105.1469
20.5411 | -105.1429
20.5449 | -105.1395
20.5486 | -105.1353
20.5523 | -105.1309
20.556 | -105.127
20.5597 | -105.1223
20.5635 | -105.1157
20.5672 | -105.1105
20.5709 | -105.1042
20.5746 | -105.0985
20.5783 | -105.0946
20.5821 | -105.0911
20.5858 | -105.0844
20.5895 | -105.0759
20.5932 | -105.0685
20.5969 | -105.0611
20.6007 | -105.0544
20.6044 | -105.0466
20.6081 | -105.0402
20.6118 | -105.0327
20.6155 | -105.0271
20.6192 | -105.0228
20.623 | -105.0171
20.6267 | -105.0111
20.6304 | -105.0055
20.6341 | -105.0003
20.6378 | -104.9922
20.6416 | -104.9858
20.6453 | -104.9796
20.649 | -104.9729
20.6527 | -104.9655
20.6564 | -104.9611
20.6602 | -104.9533
20.6639 | -104.9465
20.6676 | -104.9411
20.6713 | -104.9361
20.675 | -104.9336
20.6788 | -104.9298
20.6825 | -104.9227
20.6862 | -104.9175
20.6899 | -104.9106
20.6936 | -104.9016
20.6974 | -104.8953
20.7011 | -104.8893
20.7048 | -104.8835
20.7085 | -104.8769
20.7122 | -104.8709
20.7159 | -104.864
20.7197 | -104.8568
20.7234 | -104.8495
20.7271 | -104.839
20.7308 | -104.829
20.7345 | -104.8189
20.7383 | -104.8077
20.742 | -104.7978
20.7457 | -104.7883
20.7494 | -104.7772
20.7531 | -104.7652
20.7569 | -104.7528
20.7606 | -104.7418
20.7643 | -104.7322
20.768 | -104.7227
20.7717 | -104.7108
20.7755 | -104.6983
20.7792 | -104.6827
20.7829 | -104.6703
20.7866 | -104.6585
20.7903 | -104.6447
20.794 | -104.6297
20.7978 | -104.6152
20.8015 | -104.6038
20.8052 | -104.5923
20.8089 | -104.5778
20.8126 | -104.5641
20.8164 | -104.5536
20.8201 | -104.5442
20.8238 | -104.5332
20.8275 | -104.5215
20.8312 | -104.5091
20.835 | -104.4988
20.8387 | -104.4898
20.8424 | -104.4774
20.8461 | -104.4669
20.8498 | -104.4593
20.8536 | -104.4524
20.8573 | -104.4471
20.861 | -104.4426
20.8647 | -104.4393
20.8684 | -104.4363
20.8722 | -104.4343
20.8759 | -104.4348
20.8796 | -104.4356
20.8833 | -104.437
20.887 | -104.4399
20.8907 | -104.4434
20.8945 | -104.4474
20.8982 | -104.4516
20.9019 | -104.4554
20.9056 | -104.4612
20.9093 | -104.469
20.9131 | -104.4776
20.9168 | -104.486
20.9205 | -104.4946
20.9242 | -104.5081
20.9279 | -104.5176
20.9317 | -104.5277
20.9354 | -104.5389
20.9391 | -104.5505
20.9428 | -104.5625
20.9465 | -104.5758
20.9503 | -104.5884
20.954 | -104.5988
20.9577 | -104.6114
20.9614 | -104.6252
20.9651 | -104.6383
20.9689 | -104.6512
20.9726 | -104.6613
20.9763 | -104.6722
20.98 | -104.6846
20.9837 | -104.6948
20.9874 | -104.7076
20.9912 | -104.7223
20.9949 | -104.7343
20.9986 | -104.7457
21.0023 | -104.7572
21.006 | -104.7695
21.0098 | -104.7824
21.0135 | -104.7961
21.0172 | -104.8105
21.0209 | -104.8248
21.0246 | -104.8426
21.0284 | -104.8573
21.0321 | -104.8726
21.0358 | -104.8895
21.0395 | -104.9036
21.0432 | -104.9181
21.047 | -104.9349
21.0507 | -104.9511
21.0544 | -104.9675
21.0581 | -104.9824
21.0618 | -104.9967
21.0656 | -105.0142
21.0693 | -105.0338
21.073 | -105.0532
21.0767 | -105.0733
21.0804 | -105.0911
21.0841 | -105.1087
21.0879 | -105.1275
21.0916 | -105.1487
21.0953 | -105.171
21.099 | -105.1907
21.1027 | -105.2111
21.1065 | -105.2296
21.1102 | -105.2495
21.1139 | -105.2675
21.1176 | -105.2854
21.1213 | -105.3041
21.1251 | -105.3201
21.1288 | -105.3379
21.1325 | -105.3543
21.1362 | -105.3674
21.1399 | -105.3845
21.1437 | -105.3995
21.1474 | -105.4092
21.1511 | -105.4185
21.1548 | -105.4269
21.1585 | -105.435
21.1623 | -105.4451
21.166 | -105.4536
21.1697 | -105.4599
21.1734 | -105.4627
21.1771 | -105.4677
21.1808 | -105.4731
21.1846 | -105.4764
21.1883 | -105.4784
21.192 | -105.4784
21.1957 | -105.4768
21.1994 | -105.4739
21.2032 | -105.4709
21.2069 | -105.4664
21.2106 | -105.458
21.2143 | -105.4504
21.218 | -105.442
21.2218 | -105.4331
21.2255 | -105.4231
21.2292 | -105.412
21.2329 | -105.4001
21.2366 | -105.3893
21.2404 | -105.3788
21.2441 | -105.3684
21.2478 | -105.3585
21.2515 | -105.3487
21.2552 | -105.3391
21.2589 | -105.3281
21.2627 | -105.316
21.2664 | -105.3061
21.2701 | -105.2966
21.2738 | -105.2899
21.2775 | -105.2822
21.2813 | -105.2743
21.285 | -105.2677
21.2887 | -105.2605
21.2924 | -105.2536
21.2961 | -105.2494
21.2999 | -105.2461
21.3036 | -105.245
21.3073 | -105.244
21.311 | -105.2423
21.3147 | -105.2421
21.3185 | -105.2407
21.3222 | -105.2403
21.3259 | -105.2391
21.3296 | -105.2395
21.3333 | -105.2404
21.3371 | -105.2431
21.3408 | -105.2467
21.3445 | -105.2516
21.3482 | -105.2562
21.3519 | -105.2628
21.3556 | -105.2711
21.3594 | -105.2797
21.3631 | -105.2904
21.3668 | -105.2976
21.3705 | -105.3062
21.3742 | -105.316
21.378 | -105.3251
21.3817 | -105.3352
21.3854 | -105.3459
21.3891 | -105.3563
21.3928 | -105.3666
21.3966 | -105.3765
21.4003 | -105.3851
21.404 | -105.3954
21.4077 | -105.4068
21.4114 | -105.4149
21.4152 | -105.4214
21.4189 | -105.428
21.4226 | -105.4341
21.4263 | -105.4406
21.43 | -105.4464
21.4338 | -105.4519
21.4375 | -105.4594
21.4412 | -105.4646
21.4449 | -105.4693
21.4486 | -105.4751
21.4523 | -105.4792
21.4561 | -105.4798
21.4598 | -105.4738
21.4635 | -105.4724
21.4672 | -105.4699
21.4709 | -105.4668
21.4747 | -105.4643
21.4784 | -105.4625
21.4821 | -105.4569
21.4858 | -105.4512
21.4895 | -105.4454
21.4933 | -105.4397
21.497 | -105.4326
21.5007 | -105.4245
21.5044 | -105.4196
21.5081 | -105.414
21.5119 | -105.4068
21.5156 | -105.4006
21.5193 | -105.3937
21.523 | -105.387
21.5267 | -105.3775
21.5305 | -105.37
21.5342 | -105.3617
21.5379 | -105.3525
21.5416 | -105.3433
21.5453 | -105.3353
21.549 | -105.3265
21.5528 | -105.3192
21.5565 | -105.3127
21.5602 | -105.3084
21.5639 | -105.3032
21.5676 | -105.2967
21.5714 | -105.2885
21.5751 | -105.2804
21.5788 | -105.2706
21.5825 | -105.2634
21.5862 | -105.2559
21.59 | -105.2521
21.5937 | -105.2484
21.5974 | -105.242
21.6011 | -105.2342
21.6048 | -105.2255
21.6086 | -105.219
21.6123 | -105.2135
21.616 | -105.2072
21.6197 | -105.2006
21.6234 | -105.1943
21.6272 | -105.1892
21.6309 | -105.1845
21.6346 | -105.1781
21.6383 | -105.1696
21.642 | -105.1607
21.6457 | -105.1499
21.6495 | -105.1385
21.6532 | -105.128
21.6569 | -105.1179
21.6606 | -105.1068
21.6643 | -105.0948
21.6681 | -105.0808
21.6718 | -105.0672
21.6755 | -105.0545
21.6792 | -105.0422
21.6829 | -105.0295
21.6867 | -105.0169
21.6904 | -105.0026
21.6941 | -104.9887
21.6978 | -104.9738
21.7015 | -104.9582
21.7053 | -104.9433
21.709 | -104.927
21.7127 | -104.9118
21.7164 | -104.8965
21.7201 | -104.8819
21.7238 | -104.8674
21.7276 | -104.8523
21.7313 | -104.837
21.735 | -104.8201
21.7387 | -104.804
21.7424 | -104.7877
21.7462 | -104.7712
21.7499 | -104.7552
21.7536 | -104.7405
21.7573 | -104.7275
21.761 | -104.7154
21.7648 | -104.7035
21.7685 | -104.6908
21.7722 | -104.6786
21.7759 | -104.6645
21.7796 | -104.6504
21.7834 | -104.6375
21.7871 | -104.6244
21.7908 | -104.6119
21.7945 | -104.6005
21.7982 | -104.5896
21.802 | -104.5789
21.8057 | -104.5684
21.8094 | -104.5602
21.8131 | -104.5509
21.8168 | -104.5406
21.8205 | -104.5313
21.8243 | -104.5201
21.828 | -104.5109
21.8317 | -104.5015
21.8354 | -104.4928
21.8391 | -104.4846
21.8429 | -104.477
21.8466 | -104.4706
21.8503 | -104.4634
21.854 | -104.4532
21.8577 | -104.4441
21.8615 | -104.4346
21.8652 | -104.4261
21.8689 | -104.4183
21.8726 | -104.4113
21.8763 | -104.4046
21.8801 | -104.3999
21.8838 | -104.3954
21.8875 | -104.3899
21.8912 | -104.3813
21.8949 | -104.3751
21.8987 | -104.3661
21.9024 | -104.3591
21.9061 | -104.3544
21.9098 | -104.3495
21.9135 | -104.3466
21.9172 | -104.341
21.921 | -104.3363
21.9247 | -104.3337
21.9284 | -104.3298
21.9321 | -104.3259
21.9358 | -104.323
21.9396 | -104.3221
21.9433 | -104.3226
21.947 | -104.3239
21.9507 | -104.3247
21.9544 | -104.3276
21.9582 | -104.3332
21.9619 | -104.3349
21.9656 | -104.3377
21.9693 | -104.3418
21.973 | -104.3458
21.9768 | -104.3504
21.9805 | -104.356
21.9842 | -104.3615
21.9879 | -104.3657
21.9916 | -104.3715
21.9954 | -104.3779
21.9991 | -104.3847
22.0028 | -104.391
22.0065 | -104.3986
22.0102 | -104.4075
22.0139 | -104.4161
22.0177 | -104.4249
22.0214 | -104.4326
22.0251 | -104.4387
22.0288 | -104.4446
22.0325 | -104.4497
22.0363 | -104.4541
22.04 | -104.4593
22.0437 | -104.4645
22.0474 | -104.467
22.0511 | -104.4693
22.0549 | -104.4727
22.0586 | -104.4768
22.0623 | -104.4796
22.066 | -104.4824
22.0697 | -104.4843
22.0735 | -104.4857
22.0772 | -104.4871
22.0809 | -104.4894
22.0846 | -104.4912
22.0883 | -104.4917
22.0921 | -104.4947
22.0958 | -104.4963
22.0995 | -104.4986
22.1032 | -104.5019
22.1069 | -104.5035
22.1106 | -104.5038
22.1144 | -104.5026
22.1181 | -104.5035
22.1218 | -104.5039
22.1255 | -104.5019
22.1292 | -104.5032
22.133 | -104.5038
22.1367 | -104.5065
22.1404 | -104.5066
22.1441 | -104.5067
22.1478 | -104.5062
22.1516 | -104.5056
22.1553 | -104.5073
22.159 | -104.5075
22.1627 | -104.5076
22.1664 | -104.51
22.1702 | -104.5124
22.1739 | -104.5149
22.1776 | -104.5186
22.1813 | -104.5207
22.185 | -104.5218
22.1887 | -104.5241
22.1925 | -104.5256
22.1962 | -104.5277
22.1999 | -104.5314
22.2036 | -104.5347
22.2073 | -104.5358
22.2111 | -104.5387
22.2148 | -104.5401
22.2185 | -104.5422
22.2222 | -104.5451
22.2259 | -104.5464
22.2297 | -104.5469
22.2334 | -104.5474
22.2371 | -104.5484
22.2408 | -104.5487
22.2445 | -104.5507
22.2483 | -104.5531
22.252 | -104.5521
22.2557 | -104.551
22.2594 | -104.5487
22.2631 | -104.5476
22.2669 | -104.5472
22.2706 | -104.5452
22.2743 | -104.5465
22.278 | -104.5465
22.2817 | -104.5462
22.2854 | -104.5477
22.2892 | -104.5497
22.2929 | -104.5514
22.2966 | -104.5513
22.3003 | -104.553
22.304 | -104.5543
22.3078 | -104.5556
22.3115 | -104.5571
22.3152 | -104.5577
22.3189 | -104.5573
22.3226 | -104.5558
22.3264 | -104.5557
22.3301 | -104.5569
22.3338 | -104.5576
22.3375 | -104.5608
22.3412 | -104.5655
22.345 | -104.5699
22.3487 | -104.5738
22.3524 | -104.5762
22.3561 | -104.5768
22.3598 | -104.578
22.3636 | -104.5834
22.3673 | -104.5873
22.371 | -104.5898
22.3747 | -104.5909
22.3784 | -104.5923
22.3821 | -104.5928
22.3859 | -104.5923
22.3896 | -104.5925
22.3933 | -104.5923
22.397 | -104.5904
22.4007 | -104.5897
22.4045 | -104.5907
22.4082 | -104.5873
22.4119 | -104.5846
22.4156 | -104.5813
22.4193 | -104.5749
22.4231 | -104.5679
22.4268 | -104.5612
22.4305 | -104.5567
22.4342 | -104.5507
22.4379 | -104.5423
22.4417 | -104.5343
22.4454 | -104.5245
22.4491 | -104.5125
22.4528 | -104.4987
22.4565 | -104.4862
22.4603 | -104.4728
22.464 | -104.4584
22.4677 | -104.4421
22.4714 | -104.4214
22.4751 | -104.3997
22.4788 | -104.375
22.4826 | -104.3525
22.4863 | -104.3305
22.49 | -104.3054
22.4937 | -104.2799
22.4974 | -104.2535
22.5012 | -104.2267
22.5049 | -104.1979
22.5086 | -104.1686
22.5123 | -104.1399
22.516 | -104.1106
22.5198 | -104.0821
22.5235 | -104.0532
22.5272 | -104.0231
22.5309 | -103.9911
22.5346 | -103.9517
22.5384 | -103.9205
22.5421 | -103.8887
22.5458 | -103.8566
22.5495 | -103.8251
22.5532 | -103.7935
22.557 | -103.7607
22.5607 | -103.7301
22.5644 | -103.7003
22.5681 | -103.6701
22.5718 | -103.6399
22.5755 | -103.61
22.5793 | -103.5825
22.583 | -103.557
22.5867 | -103.532
22.5904 | -103.5082
22.5941 | -103.4837
22.5979 | -103.4611
22.6016 | -103.4412
22.6053 | -103.4221
22.609 | -103.4055
22.6127 | -103.3905
22.6165 | -103.3787
22.6202 | -103.3687
22.6239 | -103.3599
22.6276 | -103.3544
22.6313 | -103.3506
22.6351 | -103.3489
22.6388 | -103.3493
22.6425 | -103.3516
22.6462 | -103.3549
22.6499 | -103.3588
22.6536 | -103.3646
22.6574 | -103.3724
22.6611 | -103.3832
22.6648 | -103.3958
22.6685 | -103.4097
22.6722 | -103.4259
22.676 | -103.4456
22.6797 | -103.4658
22.6834 | -103.4851
22.6871 | -103.5071
22.6908 | -103.5308
22.6946 | -103.5566
22.6983 | -103.5825
22.702 | -103.6065
22.7057 | -103.6326
22.7094 | -103.6601
22.7132 | -103.6879
22.7169 | -103.7142
22.7206 | -103.7388
22.7243 | -103.7639
22.728 | -103.7889
22.7318 | -103.8131
22.7355 | -103.8356
22.7392 | -103.8593
22.7429 | -103.8822
22.7466 | -103.9028
22.7503 | -103.9216
22.7541 | -103.9389
22.7578 | -103.9551
22.7615 | -103.9702
22.7652 | -103.9835
22.7689 | -103.9988
22.7727 | -104.0131
22.7764 | -104.0243
22.7801 | -104.0351
22.7838 | -104.0452
22.7875 | -104.0556
22.7913 | -104.0647
22.795 | -104.0744
22.7987 | -104.0859
22.8024 | -104.0919
22.8061 | -104.0952
22.8099 | -104.1006
22.8136 | -104.1054
22.8173 | -104.1104
22.821 | -104.1172
22.8247 | -104.1242
22.8285 | -104.1299
22.8322 | -104.1346
22.8359 | -104.1401
22.8396 | -104.1441
22.8433 | -104.1495
22.847 | -104.1558
22.8508 | -104.1635
22.8545 | -104.1731
22.8582 | -104.1819
22.8619 | -104.1896
22.8656 | -104.1971
22.8694 | -104.2051
22.8731 | -104.2139
22.8768 | -104.2231
22.8805 | -104.2329
22.8842 | -104.2421
22.888 | -104.2523
22.8917 | -104.2617
22.8954 | -104.2733
22.8991 | -104.2859
22.9028 | -104.2974
22.9066 | -104.3072
22.9103 | -104.3175
22.914 | -104.3267
22.9177 | -104.3372
22.9214 | -104.3488
22.9252 | -104.3578
22.9289 | -104.3682
22.9326 | -104.3787
22.9363 | -104.387
22.94 | -104.3943
22.9437 | -104.4023
22.9475 | -104.4109
22.9512 | -104.4174
22.9549 | -104.4233
22.9586 | -104.4292
22.9623 | -104.4332
22.9661 | -104.4362
22.9698 | -104.4371
22.9735 | -104.4386
22.9772 | -104.4393
22.9809 | -104.4378
22.9847 | -104.4351
22.9884 | -104.4294
22.9921 | -104.4236
22.9958 | -104.4188
22.9995 | -104.4115
23.0033 | -104.4043
23.007 | -104.3992
23.0107 | -104.3915
23.0144 | -104.3824
23.0181 | -104.3719
23.0219 | -104.361
23.0256 | -104.3518
23.0293 | -104.3397
23.033 | -104.3239
23.0367 | -104.3088
23.0404 | -104.2975
23.0442 | -104.2856
23.0479 | -104.2725
23.0516 | -104.2598
23.0553 | -104.2458
23.059 | -104.2324
23.0628 | -104.2178
23.0665 | -104.2037
23.0702 | -104.1919
23.0739 | -104.1791
23.0776 | -104.1657
23.0814 | -104.1532
23.0851 | -104.1437
23.0888 | -104.1336
23.0925 | -104.1231
23.0962 | -104.1115
23.1 | -104.1004
23.1037 | -104.0902
23.1074 | -104.0814
23.1111 | -104.0725
23.1148 | -104.064
23.1185 | -104.0573
23.1223 | -104.0514
23.126 | -104.0443
23.1297 | -104.0385
23.1334 | -104.0352
23.1371 | -104.0315
23.1409 | -104.0274
23.1446 | -104.0253
23.1483 | -104.0225
23.152 | -104.0216
23.1557 | -104.0225
23.1595 | -104.0229
23.1632 | -104.0248
23.1669 | -104.0291
23.1706 | -104.0319
23.1743 | -104.0346
23.1781 | -104.0397
23.1818 | -104.0469
23.1855 | -104.0545
23.1892 | -104.0605
23.1929 | -104.0691
23.1967 | -104.0789
23.2004 | -104.0884
23.2041 | -104.0984
23.2078 | -104.1095
23.2115 | -104.1216
23.2152 | -104.1319
23.219 | -104.1436
23.2227 | -104.158
23.2264 | -104.172
23.2301 | -104.1844
23.2338 | -104.1961
23.2376 | -104.2105
23.2413 | -104.2267
23.245 | -104.2432
23.2487 | -104.2579
23.2524 | -104.2745
23.2562 | -104.2907
23.2599 | -104.3063
23.2636 | -104.3248
23.2673 | -104.3426
23.271 | -104.3606
23.2748 | -104.3783
23.2785 | -104.3946
23.2822 | -104.4118
23.2859 | -104.4306
23.2896 | -104.4504
23.2934 | -104.468
23.2971 | -104.4864
23.3008 | -104.5036
23.3045 | -104.521
23.3082 | -104.5377
23.3119 | -104.5527
23.3157 | -104.5688
23.3194 | -104.5824
23.3231 | -104.5938
23.3268 | -104.6053
23.3305 | -104.6145
23.3343 | -104.6235
23.338 | -104.6322
23.3417 | -104.6416
23.3454 | -104.6517
23.3491 | -104.6615
23.3529 | -104.6699
23.3566 | -104.679
23.3603 | -104.6883
23.364 | -104.6966
23.3677 | -104.7007
23.3715 | -104.7051
23.3752 | -104.7094
23.3789 | -104.7131
23.3826 | -104.7168
23.3863 | -104.7197
23.3901 | -104.7228
23.3938 | -104.7236
23.3975 | -104.724
23.4012 | -104.7255
23.4049 | -104.7272
23.4086 | -104.7291
23.4124 | -104.7305
23.4161 | -104.7309
23.4198 | -104.7319
23.4235 | -104.7344
23.4272 | -104.7369
23.431 | -104.7405
23.4347 | -104.744
23.4384 | -104.7452
23.4421 | -104.7468
23.4458 | -104.7486
23.4496 | -104.7513
23.4533 | -104.7529
23.457 | -104.7555
23.4607 | -104.76
23.4644 | -104.7638
23.4682 | -104.7674
23.4719 | -104.771
23.4756 | -104.7741
23.4793 | -104.7791
23.483 | -104.7841
23.4868 | -104.7889
23.4905 | -104.7947
23.4942 | -104.8001
23.4979 | -104.8048
23.5016 | -104.8102
23.5053 | -104.8171
23.5091 | -104.8249
23.5128 | -104.8305
23.5165 | -104.8364
23.5202 | -104.8447
23.5239 | -104.8526
23.5277 | -104.8631
23.5314 | -104.8744
23.5351 | -104.8846
23.5388 | -104.8947
23.5425 | -104.9057
23.5463 | -104.9197
23.55 | -104.9346
23.5537 | -104.946
23.5574 | -104.9579
23.5611 | -104.9712
23.5649 | -104.9855
23.5686 | -104.9973
23.5723 | -105.0093
23.576 | -105.0215
23.5797 | -105.0354
23.5834 | -105.0488
23.5872 | -105.0615
23.5909 | -105.0742
23.5946 | -105.0847
23.5983 | -105.0963
23.602 | -105.108
23.6058 | -105.1187
23.6095 | -105.1283
23.6132 | -105.1378
23.6169 | -105.1472
23.6206 | -105.1543
23.6244 | -105.1618
23.6281 | -105.1702
23.6318 | -105.1755
23.6355 | -105.1811
23.6392 | -105.1849
23.643 | -105.1883
23.6467 | -105.1925
23.6504 | -105.1958
23.6541 | -105.1986
23.6578 | -105.2002
23.6616 | -105.1999
23.6653 | -105.2003
23.669 | -105.1983
23.6727 | -105.1963
23.6764 | -105.1941
23.6801 | -105.19
23.6839 | -105.1852
23.6876 | -105.1816
23.6913 | -105.1777
23.695 | -105.1744
23.6987 | -105.1696
23.7025 | -105.1668
23.7062 | -105.163
23.7099 | -105.1576
23.7136 | -105.1524
23.7173 | -105.1466
23.7211 | -105.1429
23.7248 | -105.1399
23.7285 | -105.1381
23.7322 | -105.136
23.7359 | -105.1323
23.7397 | -105.1288
23.7434 | -105.1244
23.7471 | -105.119
23.7508 | -105.114
23.7545 | -105.1112
23.7583 | -105.1087
23.762 | -105.1062
23.7657 | -105.1042
23.7694 | -105.1038
23.7731 | -105.105
23.7768 | -105.1066
23.7806 | -105.1077
23.7843 | -105.1101
23.788 | -105.1132
23.7917 | -105.1163
23.7954 | -105.119
23.7992 | -105.122
23.8029 | -105.1274
23.8066 | -105.1342
23.8103 | -105.1413
23.814 | -105.1483
23.8178 | -105.1552
23.8215 | -105.162
23.8252 | -105.1704
23.8289 | -105.18
23.8326 | -105.1919
23.8364 | -105.2054
23.8401 | -105.2183
23.8438 | -105.2293
23.8475 | -105.2417
23.8512 | -105.2553
23.855 | -105.2701
23.8587 | -105.286
23.8624 | -105.3005
23.8661 | -105.3151
23.8698 | -105.331
23.8735 | -105.348
23.8773 | -105.365
23.881 | -105.3815
23.8847 | -105.3977
23.8884 | -105.4145
23.8921 | -105.4334
23.8959 | -105.454
23.8996 | -105.4744
23.9033 | -105.4947
23.907 | -105.5153
23.9107 | -105.5358
23.9145 | -105.5548
23.9182 | -105.5726
23.9219 | -105.5912
23.9256 | -105.6095
23.9293 | -105.6281
23.9331 | -105.6449
23.9368 | -105.6605
23.9405 | -105.6764
23.9442 | -105.6909
23.9479 | -105.7076
23.9517 | -105.7229
23.9554 | -105.737
23.9591 | -105.7505
23.9628 | -105.7606
23.9665 | -105.769
23.9702 | -105.7794
23.974 | -105.7884
23.9777 | -105.7965
23.9814 | -105.801
23.9851 | -105.8033
23.9888 | -105.8071
23.9926 | -105.8118
23.9963 | -105.8128
24 | -105.8127
24.0037 | -105.8117
24.0074 | -105.8086
24.0112 | -105.8044
24.0149 | -105.7987
24.0186 | -105.7927
24.0223 | -105.7847
24.026 | -105.7761
24.0298 | -105.7654
24.0335 | -105.7544
24.0372 | -105.7436
24.0409 | -105.7292
24.0446 | -105.7156
24.0483 | -105.7041
24.0521 | -105.6901
24.0558 | -105.6751
24.0595 | -105.6578
24.0632 | -105.6395
24.0669 | -105.6199
24.0707 | -105.6001
24.0744 | -105.5805
24.0781 | -105.5631
24.0818 | -105.5467
24.0855 | -105.5305
24.0893 | -105.5128
24.093 | -105.4937
24.0967 | -105.4743
24.1004 | -105.4552
24.1041 | -105.4374
24.1079 | -105.4196
24.1116 | -105.4008
24.1153 | -105.3826
24.119 | -105.3649
24.1227 | -105.3492
24.1265 | -105.3339
24.1302 | -105.3179
24.1339 | -105.3013
24.1376 | -105.2873
24.1413 | -105.2735
24.145 | -105.2661
24.1488 | -105.2537
24.1525 | -105.2426
24.1562 | -105.2302
24.1599 | -105.219
24.1636 | -105.21
24.1674 | -105.2021
24.1711 | -105.1943
24.1748 | -105.1857
24.1785 | -105.1786
24.1822 | -105.1733
24.186 | -105.1694
24.1897 | -105.1663
24.1934 | -105.1624
24.1971 | -105.1582
24.2008 | -105.1534
24.2046 | -105.151
24.2083 | -105.1507
24.212 | -105.1509
24.2157 | -105.15
24.2194 | -105.1497
24.2232 | -105.1501
24.2269 | -105.1515
24.2306 | -105.1553
24.2343 | -105.1572
24.238 | -105.1587
24.2417 | -105.1604
24.2455 | -105.1602
24.2492 | -105.1611
24.2529 | -105.1599
24.2566 | -105.1582
24.2603 | -105.1577
24.2641 | -105.1563
24.2678 | -105.154
24.2715 | -105.1503
24.2752 | -105.1435
24.2789 | -105.1352
24.2827 | -105.1268
24.2864 | -105.1182
24.2901 | -105.1092
24.2938 | -105.1015
24.2975 | -105.0937
24.3013 | -105.0855
24.305 | -105.0772
24.3087 | -105.0663
24.3124 | -105.0518
24.3161 | -105.0374
24.3199 | -105.0235
24.3236 | -105.01
24.3273 | -104.9947
24.331 | -104.9756
24.3347 | -104.9548
24.3384 | -104.937
24.3422 | -104.9179
24.3459 | -104.8969
24.3496 | -104.8766
24.3533 | -104.8568
24.357 | -104.8364
24.3608 | -104.8157
24.3645 | -104.7951
24.3682 | -104.7739
24.3719 | -104.7518
24.3756 | -104.7298
24.3794 | -104.7071
24.3831 | -104.6876
24.3868 | -104.6683
24.3905 | -104.647
24.3942 | -104.6273
24.398 | -104.6093
24.4017 | -104.5914
24.4054 | -104.575
24.4091 | -104.5606
24.4128 | -104.5473
24.4166 | -104.5359
24.4203 | -104.5236
24.424 | -104.5118
24.4277 | -104.5028
24.4314 | -104.4955
24.4351 | -104.4884
24.4389 | -104.4827
24.4426 | -104.4809
24.4463 | -104.4795
24.45 | -104.4783
24.4537 | -104.4802
24.4575 | -104.4853
24.4612 | -104.4911
24.4649 | -104.4964
24.4686 | -104.5027
24.4723 | -104.5127
24.4761 | -104.5242
24.4798 | -104.5344
24.4835 | -104.5462
24.4872 | -104.5595
24.4909 | -104.5739
24.4947 | -104.5878
24.4984 | -104.6026
24.5021 | -104.6195
24.5058 | -104.6372
24.5095 | -104.6549
24.5132 | -104.6731
24.517 | -104.6932
24.5207 | -104.7129
24.5244 | -104.7298
24.5281 | -104.7468
24.5318 | -104.7656
24.5356 | -104.7856
24.5393 | -104.8055
24.543 | -104.8237
24.5467 | -104.841
24.5504 | -104.856
24.5542 | -104.8683
24.5579 | -104.8825
24.5616 | -104.899
24.5653 | -104.9118
24.569 | -104.9202
24.5728 | -104.9285
24.5765 | -104.9353
24.5802 | -104.9421
24.5839 | -104.9448
24.5876 | -104.9467
24.5914 | -104.9419
24.5951 | -104.9217
24.5988 | -104.8978
24.6025 | -104.8793
24.6062 | -104.8531
24.6099 | -104.8251
24.6137 | -104.7986
24.6174 | -104.7665
24.6211 | -104.7316
24.6248 | -104.6953
24.6285 | -104.6536
24.6323 | -104.6081
24.636 | -104.559
24.6397 | -104.5083
24.6434 | -104.454
24.6471 | -104.3975
24.6509 | -104.3378
24.6546 | -104.277
24.6583 | -104.2143
24.662 | -104.1488
24.6657 | -104.081
24.6695 | -104.0111
24.6732 | -103.9408
24.6769 | -103.8705
24.6806 | -103.783
24.6843 | -103.713
24.6881 | -103.6432
24.6918 | -103.5732
24.6955 | -103.5032
24.6992 | -103.433
24.7029 | -103.3636
24.7066 | -103.2952
24.7104 | -103.2302
24.7141 | -103.1671
24.7178 | -103.1057
24.7215 | -103.0478
24.7252 | -102.9902
24.729 | -102.9351
24.7327 | -102.8843
24.7364 | -102.8362
24.7401 | -102.7923
24.7438 | -102.7522
24.7476 | -102.7154
24.7513 | -102.684
24.755 | -102.6563
24.7587 | -102.632
24.7624 | -102.6111
24.7662 | -102.5942
24.7699 | -102.5827
24.7736 | -102.5757
24.7773 | -102.5725
24.781 | -102.5737
24.7848 | -102.58
24.7885 | -102.5901
24.7922 | -102.6029
24.7959 | -102.6203
24.7996 | -102.6428
24.8033 | -102.6689
24.8071 | -102.698
24.8108 | -102.7321
24.8145 | -102.7711
24.8182 | -102.8134
24.8219 | -102.857
24.8257 | -102.9039
24.8294 | -102.955
24.8331 | -103.0089
24.8368 | -103.0657
24.8405 | -103.125
24.8443 | -103.1869
24.848 | -103.2505
24.8517 | -103.3154
24.8554 | -103.3814
24.8591 | -103.4492
24.8629 | -103.5177
24.8666 | -103.5852
24.8703 | -103.654
24.874 | -103.7252
24.8777 | -103.7958
24.8815 | -103.8643
24.8852 | -103.932
24.8889 | -104.0005
24.8926 | -104.0683
24.8963 | -104.1341
24.9 | -104.199
24.9038 | -104.2632
24.9075 | -104.3233
24.9112 | -104.3807
24.9149 | -104.4385
24.9186 | -104.4956
24.9224 | -104.551
24.9261 | -104.6041
24.9298 | -104.6529
24.9335 | -104.7007
24.9372 | -104.7447
24.941 | -104.7931
24.9447 | -104.8493
24.9484 | -104.9025
24.9521 | -104.9429
24.9558 | -104.9822
24.9596 | -105.0156
24.9633 | -105.0419
24.967 | -105.071
24.9707 | -105.0987
24.9744 | -105.1206
24.9781 | -105.1396
24.9819 | -105.1575
24.9856 | -105.1722
24.9893 | -105.183
24.993 | -105.1944
24.9967 | -105.2031
25.0005 | -105.2128
25.0042 | -105.2216
25.0079 | -105.2303
25.0116 | -105.2366
25.0153 | -105.2397
25.0191 | -105.2445
25.0228 | -105.2469
25.0265 | -105.2502
25.0302 | -105.2551
25.0339 | -105.2593
25.0377 | -105.2649
25.0414 | -105.2667
25.0451 | -105.2698
25.0488 | -105.2739
25.0525 | -105.2762
25.0563 | -105.2783
25.06 | -105.2801
25.0637 | -105.2789
25.0674 | -105.2788
25.0711 | -105.2791
25.0748 | -105.2786
25.0786 | -105.2784
25.0823 | -105.278
25.086 | -105.2741
25.0897 | -105.2694
25.0934 | -105.2634
25.0972 | -105.2589
25.1009 | -105.2546
25.1046 | -105.2487
25.1083 | -105.2419
25.112 | -105.2363
25.1158 | -105.2322
25.1195 | -105.2274
25.1232 | -105.2223
25.1269 | -105.2182
25.1306 | -105.2139
25.1344 | -105.2095
25.1381 | -105.2054
25.1418 | -105.202
25.1455 | -105.1979
25.1492 | -105.1936
25.153 | -105.1891
25.1567 | -105.1845
25.1604 | -105.1796
25.1641 | -105.1722
25.1678 | -105.1647
25.1715 | -105.159
25.1753 | -105.1562
25.179 | -105.1534
25.1827 | -105.1511
25.1864 | -105.1491
25.1901 | -105.1459
25.1939 | -105.1423
25.1976 | -105.1405
25.2013 | -105.1387
25.205 | -105.1377
25.2087 | -105.1366
25.2125 | -105.1357
25.2162 | -105.1406
25.2199 | -105.141
25.2236 | -105.143
25.2273 | -105.1452
25.2311 | -105.1475
25.2348 | -105.1505
25.2385 | -105.1552
25.2422 | -105.1607
25.2459 | -105.1653
25.2497 | -105.1695
25.2534 | -105.1762
25.2571 | -105.1842
25.2608 | -105.191
25.2645 | -105.1986
25.2682 | -105.2054
25.272 | -105.2124
25.2757 | -105.2221
25.2794 | -105.2307
25.2831 | -105.2393
25.2868 | -105.2489
25.2906 | -105.2583
25.2943 | -105.2678
25.298 | -105.2756
25.3017 | -105.2826
25.3054 | -105.2901
25.3092 | -105.2975
25.3129 | -105.3074
25.3166 | -105.3173
25.3203 | -105.3269
25.324 | -105.3343
25.3278 | -105.3428
25.3315 | -105.3531
25.3352 | -105.3636
25.3389 | -105.3714
25.3426 | -105.3798
25.3464 | -105.3887
25.3501 | -105.3957
25.3538 | -105.4042
25.3575 | -105.4104
25.3612 | -105.4168
25.3649 | -105.4243
25.3687 | -105.4322
25.3724 | -105.4408
25.3761 | -105.4494
25.3798 | -105.4591
25.3835 | -105.4668
25.3873 | -105.4722
25.391 | -105.4774
25.3947 | -105.4809
25.3984 | -105.4863
25.4021 | -105.4919
25.4059 | -105.4968
25.4096 | -105.5043
25.4133 | -105.5109
25.417 | -105.516
25.4207 | -105.5206
25.4245 | -105.5257
25.4282 | -105.5306
25.4319 | -105.5344
25.4356 | -105.5393
25.4393 | -105.5451
25.443 | -105.5517
25.4468 | -105.5592
25.4505 | -105.5661
25.4542 | -105.5698
25.4579 | -105.5728
25.4616 | -105.576
25.4654 | -105.5812
25.4691 | -105.5865
25.4728 | -105.5884
25.4765 | -105.5902
25.4802 | -105.5928
25.484 | -105.5947
25.4877 | -105.5969
25.4914 | -105.5993
25.4951 | -105.601
25.4988 | -105.6034
25.5026 | -105.6068
25.5063 | -105.6105
25.51 | -105.6126
25.5137 | -105.6135
25.5174 | -105.615
25.5212 | -105.6186
25.5249 | -105.6219
25.5286 | -105.6235
25.5323 | -105.624
25.536 | -105.625
25.5397 | -105.625
25.5435 | -105.6228
25.5472 | -105.6216
25.5509 | -105.622
25.5546 | -105.6211
25.5583 | -105.6198
25.5621 | -105.618
25.5658 | -105.6156
25.5695 | -105.6128
25.5732 | -105.61
25.5769 | -105.6071
25.5807 | -105.6038
25.5844 | -105.6017
25.5881 | -105.5983
25.5918 | -105.593
25.5955 | -105.5873
25.5993 | -105.5836
25.603 | -105.5793
25.6067 | -105.5755
25.6104 | -105.5738
25.6141 | -105.5728
25.6179 | -105.5696
25.6216 | -105.5675
25.6253 | -105.566
25.629 | -105.5652
25.6327 | -105.5653
25.6364 | -105.5634
25.6402 | -105.5617
25.6439 | -105.5598
25.6476 | -105.5582
25.6513 | -105.5597
25.655 | -105.5616
25.6588 | -105.5639
25.6625 | -105.5664
25.6662 | -105.5697
25.6699 | -105.5737
25.6736 | -105.5764
25.6774 | -105.5781
25.6811 | -105.5798
25.6848 | -105.584
25.6885 | -105.5884
25.6922 | -105.5915
25.696 | -105.5963
25.6997 | -105.6015
25.7034 | -105.6052
25.7071 | -105.6093
25.7108 | -105.6146
25.7146 | -105.6193
25.7183 | -105.6267
25.722 | -105.6343
25.7257 | -105.6403
25.7294 | -105.6473
25.7331 | -105.6559
25.7369 | -105.6622
25.7406 | -105.6691
25.7443 | -105.6759
25.748 | -105.6834
25.7517 | -105.6944
25.7555 | -105.7041
25.7592 | -105.7138
25.7629 | -105.7227
25.7666 | -105.732
25.7703 | -105.7405
25.7741 | -105.7481
25.7778 | -105.7569
25.7815 | -105.7679
25.7852 | -105.7789
25.7889 | -105.7888
25.7927 | -105.7985
25.7964 | -105.8093
25.8001 | -105.8217
25.8038 | -105.8331
25.8075 | -105.8437
25.8113 | -105.8539
25.815 | -105.8649
25.8187 | -105.8759
25.8224 | -105.8888
25.8261 | -105.9007
25.8298 | -105.911
25.8336 | -105.9206
25.8373 | -105.9334
25.841 | -105.9464
25.8447 | -105.9569
25.8484 | -105.9672
25.8522 | -105.976
25.8559 | -105.9849
25.8596 | -105.9947
25.8633 | -106.0035
25.867 | -106.0123
25.8708 | -106.022
25.8745 | -106.0321
25.8782 | -106.0427
25.8819 | -106.0529
25.8856 | -106.0619
25.8894 | -106.0702
25.8931 | -106.079
25.8968 | -106.0892
25.9005 | -106.097
25.9042 | -106.1025
25.9079 | -106.1091
25.9117 | -106.1167
25.9154 | -106.1227
25.9191 | -106.1292
25.9228 | -106.1355
25.9265 | -106.1417
25.9303 | -106.1484
25.934 | -106.1546
25.9377 | -106.1606
25.9414 | -106.1656
25.9451 | -106.1703
25.9489 | -106.1736
25.9526 | -106.1784
25.9563 | -106.1844
25.96 | -106.1913
25.9637 | -106.1955
25.9675 | -106.1994
25.9712 | -106.2031
25.9749 | -106.2062
25.9786 | -106.2076
25.9823 | -106.2095
25.9861 | -106.212
25.9898 | -106.2139
25.9935 | -106.2159
25.9972 | -106.2181
26.0009 | -106.2214
26.0046 | -106.2262
26.0084 | -106.23
26.0121 | -106.234
26.0158 | -106.2359
26.0195 | -106.2388
26.0232 | -106.2412
26.027 | -106.2437
26.0307 | -106.2462
26.0344 | -106.2485
26.0381 | -106.2494
26.0418 | -106.2517
26.0456 | -106.2541
26.0493 | -106.2549
26.053 | -106.2549
26.0567 | -106.2551
26.0604 | -106.2561
26.0642 | -106.2604
26.0679 | -106.2651
26.0716 | -106.2662
26.0753 | -106.266
26.079 | -106.2674
26.0828 | -106.2676
26.0865 | -106.2681
26.0902 | -106.2702
26.0939 | -106.2701
26.0976 | -106.27
26.1013 | -106.2705
26.1051 | -106.2704
26.1088 | -106.2702
26.1125 | -106.2702
26.1162 | -106.269
26.1199 | -106.2673
26.1237 | -106.2658
26.1274 | -106.2652
26.1311 | -106.2645
26.1348 | -106.2626
26.1385 | -106.2619
26.1423 | -106.2618
26.146 | -106.26
26.1497 | -106.2582
26.1534 | -106.2569
26.1571 | -106.2556
26.1609 | -106.2538
26.1646 | -106.2518
26.1683 | -106.2477
26.172 | -106.2422
26.1757 | -106.239
26.1795 | -106.2371
26.1832 | -106.2343
26.1869 | -106.232
26.1906 | -106.2279
26.1943 | -106.2232
26.198 | -106.2173
26.2018 | -106.2123
26.2055 | -106.208
26.2092 | -106.2046
26.2129 | -106.2023
26.2166 | -106.1988
26.2204 | -106.1952
26.2241 | -106.1914
26.2278 | -106.1884
26.2315 | -106.1848
26.2352 | -106.1799
26.239 | -106.1735
26.2427 | -106.1669
26.2464 | -106.1603
26.2501 | -106.1535
26.2538 | -106.1457
26.2576 | -106.138
26.2613 | -106.1324
26.265 | -106.1256
26.2687 | -106.1192
26.2724 | -106.1145
26.2762 | -106.1099
26.2799 | -106.1042
26.2836 | -106.0973
26.2873 | -106.0907
26.291 | -106.0825
26.2947 | -106.0775
26.2985 | -106.0722
26.3022 | -106.0671
26.3059 | -106.0618
26.3096 | -106.0563
26.3133 | -106.0498
26.3171 | -106.0445
26.3208 | -106.0381
26.3245 | -106.0323
26.3282 | -106.027
26.3319 | -106.0197
26.3357 | -106.0132
26.3394 | -106.0075
26.3431 | -106.0024
26.3468 | -105.9969
26.3505 | -105.9921
26.3543 | -105.9872
26.358 | -105.9823
26.3617 | -105.979
26.3654 | -105.9762
26.3691 | -105.9724
26.3728 | -105.9694
26.3766 | -105.964
26.3803 | -105.9612
26.384 | -105.9593
26.3877 | -105.9588
26.3914 | -105.9589
26.3952 | -105.958
26.3989 | -105.9567
26.4026 | -105.9562
26.4063 | -105.9565
26.41 | -105.9568
26.4138 | -105.957
26.4175 | -105.9579
26.4212 | -105.9573
26.4249 | -105.9577
26.4286 | -105.9597
26.4324 | -105.9604
26.4361 | -105.959
26.4398 | -105.9591
26.4435 | -105.9623
26.4472 | -105.9638
26.451 | -105.9635
26.4547 | -105.9636
26.4584 | -105.9638
26.4621 | -105.9636
26.4658 | -105.9639
26.4695 | -105.9634
26.4733 | -105.9633
26.477 | -105.9642
26.4807 | -105.9655
26.4844 | -105.9656
26.4881 | -105.9646
26.4919 | -105.9639
26.4956 | -105.9633
26.4993 | -105.9614
26.503 | -105.9586
26.5067 | -105.956
26.5105 | -105.9536
26.5142 | -105.9518
26.5179 | -105.9508
26.5216 | -105.9479
26.5253 | -105.9438
26.5291 | -105.9403
26.5328 | -105.9374
26.5365 | -105.9348
26.5402 | -105.9321
26.5439 | -105.9302
26.5477 | -105.9264
26.5514 | -105.9216
26.5551 | -105.9167
26.5588 | -105.9118
26.5625 | -105.9062
26.5662 | -105.9015
26.57 | -105.8987
26.5737 | -105.8957
26.5774 | -105.8904
26.5811 | -105.8835
26.5848 | -105.8779
26.5886 | -105.8732
26.5923 | -105.8694
26.596 | -105.8644
26.5997 | -105.858
26.6034 | -105.8517
26.6072 | -105.8467
26.6109 | -105.8424
26.6146 | -105.8381
26.6183 | -105.8341
26.622 | -105.8293
26.6258 | -105.8265
26.6295 | -105.8227
26.6332 | -105.8173
26.6369 | -105.8131
26.6406 | -105.8095
26.6444 | -105.8056
26.6481 | -105.8025
26.6518 | -105.8002
26.6555 | -105.798
26.6592 | -105.7955
26.6629 | -105.7943
26.6667 | -105.7938
26.6704 | -105.7912
26.6741 | -105.7901
26.6778 | -105.7892
26.6815 | -105.7874
26.6853 | -105.7868
26.689 | -105.7881
26.6927 | -105.7891
26.6964 | -105.7896
26.7001 | -105.7903
26.7039 | -105.7909
26.7076 | -105.792
26.7113 | -105.7947
26.715 | -105.7983
26.7187 | -105.8021
26.7225 | -105.8065
26.7262 | -105.8107
26.7299 | -105.8158
26.7336 | -105.8212
26.7373 | -105.8246
26.7411 | -105.8284
26.7448 | -105.8332
26.7485 | -105.8385
26.7522 | -105.846
26.7559 | -105.8524
26.7596 | -105.8574
26.7634 | -105.863
26.7671 | -105.8676
26.7708 | -105.8735
26.7745 | -105.8794
26.7782 | -105.8848
26.782 | -105.8905
26.7857 | -105.8983
26.7894 | -105.9057
26.7931 | -105.9123
26.7968 | -105.9201
26.8006 | -105.9275
26.8043 | -105.9325
26.808 | -105.9376
26.8117 | -105.9436
26.8154 | -105.9491
26.8192 | -105.953
26.8229 | -105.9569
26.8266 | -105.9599
26.8303 | -105.966
26.834 | -105.9701
26.8377 | -105.9728
26.8415 | -105.9758
26.8452 | -105.978
26.8489 | -105.9801
26.8526 | -105.9822
26.8563 | -105.9822
26.8601 | -105.9827
26.8638 | -105.9848
26.8675 | -105.987
26.8712 | -105.9867
26.8749 | -105.9865
26.8787 | -105.987
26.8824 | -105.9874
26.8861 | -105.9875
26.8898 | -105.9863
26.8935 | -105.984
26.8973 | -105.9808
26.901 | -105.9797
26.9047 | -105.9797
26.9084 | -105.9773
26.9121 | -105.9732
26.9159 | -105.9704
26.9196 | -105.9686
26.9233 | -105.9665
26.927 | -105.9627
26.9307 | -105.9592
26.9344 | -105.9572
26.9382 | -105.9549
26.9419 | -105.953
26.9456 | -105.9511
26.9493 | -105.9481
26.953 | -105.9424
26.9568 | -105.9393
26.9605 | -105.9373
26.9642 | -105.9333
26.9679 | -105.9308
26.9716 | -105.9282
26.9754 | -105.9248
26.9791 | -105.9225
26.9828 | -105.9218
26.9865 | -105.9211
26.9902 | -105.9196
26.994 | -105.9191
26.9977 | -105.9179
27.0014 | -105.9176
27.0051 | -105.9183
27.0088 | -105.9195
27.0126 | -105.9216
27.0163 | -105.922
27.02 | -105.9206
27.0237 | -105.9197
27.0274 | -105.9185
27.0311 | -105.9183
27.0349 | -105.9173
27.0386 | -105.9168
27.0423 | -105.9164
27.046 | -105.9161
27.0497 | -105.9177
27.0535 | -105.9185
27.0572 | -105.9205
27.0609 | -105.9241
27.0646 | -105.9267
27.0683 | -105.9287
27.0721 | -105.9313
27.0758 | -105.9362
27.0795 | -105.9404
27.0832 | -105.9444
27.0869 | -105.9485
27.0907 | -105.9511
27.0944 | -105.9524
27.0981 | -105.953
27.1018 | -105.9548
27.1055 | -105.9556
27.1093 | -105.9562
27.113 | -105.9568
27.1167 | -105.958
27.1204 | -105.9582
27.1241 | -105.9576
27.1278 | -105.9559
27.1316 | -105.9536
27.1353 | -105.9507
27.139 | -105.9474
27.1427 | -105.943
27.1464 | -105.94
27.1502 | -105.9375
27.1539 | -105.9351
27.1576 | -105.9312
27.1613 | -105.9249
27.165 | -105.9187
27.1688 | -105.9138
27.1725 | -105.9081
27.1762 | -105.9013
27.1799 | -105.894
27.1836 | -105.8875
27.1874 | -105.8812
27.1911 | -105.8755
27.1948 | -105.8699
27.1985 | -105.8639
27.2022 | -105.8594
27.206 | -105.8558
27.2097 | -105.8491
27.2134 | -105.8413
27.2171 | -105.8361
27.2208 | -105.8319
27.2245 | -105.8269
27.2283 | -105.8231
27.232 | -105.8182
27.2357 | -105.8124
27.2394 | -105.8069
27.2431 | -105.8
27.2469 | -105.7936
27.2506 | -105.7874
27.2543 | -105.7839
27.258 | -105.7801
27.2617 | -105.7757
27.2655 | -105.7718
27.2692 | -105.7691
27.2729 | -105.7658
27.2766 | -105.7624
27.2803 | -105.7598
27.2841 | -105.7549
27.2878 | -105.7507
27.2915 | -105.7486
27.2952 | -105.746
27.2989 | -105.7413
27.3026 | -105.7352
27.3064 | -105.7299
27.3101 | -105.7256
27.3138 | -105.7205
27.3175 | -105.7153
27.3212 | -105.7095
27.325 | -105.7049
27.3287 | -105.6997
27.3324 | -105.6944
27.3361 | -105.6915
27.3398 | -105.6858
27.3436 | -105.6786
27.3473 | -105.6722
27.351 | -105.6654
27.3547 | -105.6581
27.3584 | -105.653
27.3622 | -105.6464
27.3659 | -105.6395
27.3696 | -105.6332
27.3733 | -105.627
27.377 | -105.6208
27.3808 | -105.6152
27.3845 | -105.6103
27.3882 | -105.6065
27.3919 | -105.6005
27.3956 | -105.5943
27.3993 | -105.588
27.4031 | -105.5831
27.4068 | -105.58
27.4105 | -105.5762
27.4142 | -105.5736
27.4179 | -105.5719
27.4217 | -105.5712
27.4254 | -105.5701
27.4291 | -105.5687
27.4328 | -105.5691
27.4365 | -105.5688
27.4403 | -105.5689
27.444 | -105.5678
27.4477 | -105.566
27.4514 | -105.568
27.4551 | -105.5702
27.4589 | -105.5709
27.4626 | -105.5717
27.4663 | -105.5737
27.47 | -105.5756
27.4737 | -105.5789
27.4775 | -105.5837
27.4812 | -105.5878
27.4849 | -105.5917
27.4886 | -105.5962
27.4923 | -105.6007
27.496 | -105.6056
27.4998 | -105.6111
27.5035 | -105.6164
27.5072 | -105.623
27.5109 | -105.6287
27.5146 | -105.6345
27.5184 | -105.6404
27.5221 | -105.6466
27.5258 | -105.6556
27.5295 | -105.6628
27.5332 | -105.668
27.537 | -105.6734
27.5407 | -105.6787
27.5444 | -105.6845
27.5481 | -105.6919
27.5518 | -105.6974
27.5556 | -105.7006
27.5593 | -105.7054
27.563 | -105.7104
27.5667 | -105.7153
27.5704 | -105.7209
27.5742 | -105.7263
27.5779 | -105.731
27.5816 | -105.7335
27.5853 | -105.738
27.589 | -105.7412
27.5927 | -105.7431
27.5965 | -105.7457
27.6002 | -105.7491
27.6039 | -105.7512
27.6076 | -105.753
27.6113 | -105.7548
27.6151 | -105.7565
27.6188 | -105.7593
27.6225 | -105.7621
27.6262 | -105.7655
27.6299 | -105.7688
27.6337 | -105.77
27.6374 | -105.7717
27.6411 | -105.7716
27.6448 | -105.7709
27.6485 | -105.7711
27.6523 | -105.771
27.656 | -105.7699
27.6597 | -105.7706
27.6634 | -105.7694
27.6671 | -105.767
27.6709 | -105.7676
27.6746 | -105.7687
27.6783 | -105.769
27.682 | -105.7696
27.6857 | -105.7691
27.6894 | -105.7697
27.6932 | -105.7696
27.6969 | -105.7695
27.7006 | -105.7706
27.7043 | -105.771
27.708 | -105.7723
27.7118 | -105.7728
27.7155 | -105.7734
27.7192 | -105.7739
27.7229 | -105.7759
27.7266 | -105.7786
27.7304 | -105.781
27.7341 | -105.783
27.7378 | -105.7855
27.7415 | -105.7877
27.7452 | -105.7896
27.749 | -105.7927
27.7527 | -105.7958
27.7564 | -105.8002
27.7601 | -105.804
27.7638 | -105.8064
27.7675 | -105.8094
27.7713 | -105.8135
27.775 | -105.8176
27.7787 | -105.8214
27.7824 | -105.8265
27.7861 | -105.8325
27.7899 | -105.8385
27.7936 | -105.8446
27.7973 | -105.8501
27.801 | -105.8561
27.8047 | -105.8608
27.8085 | -105.8644
27.8122 | -105.8685
27.8159 | -105.8734
27.8196 | -105.8763
27.8233 | -105.8782
27.8271 | -105.8826
27.8308 | -105.8875
27.8345 | -105.894
27.8382 | -105.9003
27.8419 | -105.9035
27.8457 | -105.9081
27.8494 | -105.9142
27.8531 | -105.92
27.8568 | -105.9255
27.8605 | -105.9317
27.8642 | -105.9368
27.868 | -105.9406
27.8717 | -105.9474
27.8754 | -105.9555
27.8791 | -105.9628
27.8828 | -105.9704
27.8866 | -105.9772
27.8903 | -105.9836
27.894 | -105.9898
27.8977 | -105.9971
27.9014 | -106.0035
27.9052 | -106.0122
27.9089 | -106.0207
27.9126 | -106.0279
27.9163 | -106.0347
27.92 | -106.0387
27.9238 | -106.0441
27.9275 | -106.0507
27.9312 | -106.0571
27.9349 | -106.0635
27.9386 | -106.0697
27.9424 | -106.0738
27.9461 | -106.0782
27.9498 | -106.0833
27.9535 | -106.0879
27.9572 | -106.094
27.9609 | -106.099
27.9647 | -106.1026
27.9684 | -106.1055
27.9721 | -106.1075
27.9758 | -106.1119
27.9795 | -106.1148
27.9833 | -106.1182
27.987 | -106.1211
27.9907 | -106.1238
27.9944 | -106.1263
27.9981 | -106.1285
28.0019 | -106.1294