## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 231 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 233
Hood | 02-92B/P | 172
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1007.6 | hPa |  | PRC PRO 01
Humidity | 44.9 | %/Hr
Internal temperature of calibration bench | 23.27 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1011 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.89 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.01 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3062 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 22/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 24/05/2018
Engineer | Anthony HUE | Approval |  | 29/05/2018
CEO | Robert GLEMOT | Approval |  | 29/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3194 | 57.8075
0.012209 | -36.0029 | 38.7947
0.018313 | -35.0204 | 28.3972
0.024417 | -34.5599 | 21.9416
0.030521 | -34.3741 | 17.962
0.036626 | -34.2249 | 14.9477
0.04273 | -34.1332 | 12.7942
0.048834 | -34.0821 | 11.1479
0.054938 | -34.0493 | 9.8984
0.061043 | -34.0206 | 8.9042
0.067147 | -34.0047 | 8.0905
0.073251 | -33.9904 | 7.3372
0.079355 | -33.9799 | 6.6919
0.08546 | -33.9718 | 6.1515
0.091564 | -33.9666 | 5.7037
0.097668 | -33.964 | 5.3183
0.10377 | -33.9618 | 4.9853
0.10988 | -33.96 | 4.6657
0.11598 | -33.959 | 4.3755
0.12209 | -33.9586 | 4.1111
0.12819 | -33.9578 | 3.8684
0.13429 | -33.9558 | 3.6507
0.1404 | -33.9543 | 3.4742
0.1465 | -33.9536 | 3.2978
0.15261 | -33.9538 | 3.1334
0.15871 | -33.9547 | 2.9776
0.16482 | -33.9554 | 2.828
0.17092 | -33.9545 | 2.6951
0.17702 | -33.9536 | 2.5642
0.18313 | -33.9524 | 2.4432
0.18923 | -33.9509 | 2.3292
0.19534 | -33.9496 | 2.2224
0.20144 | -33.949 | 2.12
0.20754 | -33.9496 | 2.0267
0.21365 | -33.95 | 1.9337
0.21975 | -33.9507 | 1.846
0.22586 | -33.9516 | 1.7636
0.23196 | -33.9525 | 1.6864
0.23807 | -33.9532 | 1.6124
0.24417 | -33.9539 | 1.5425
0.25027 | -33.9546 | 1.4718
0.25638 | -33.955 | 1.4023
0.26248 | -33.9554 | 1.3348
0.26859 | -33.9556 | 1.269
0.27469 | -33.9557 | 1.2064
0.2808 | -33.9557 | 1.1532
0.2869 | -33.9559 | 1.0986
0.293 | -33.9561 | 1.0468
0.29911 | -33.9564 | 0.99814
0.30521 | -33.957 | 0.95181
0.31132 | -33.9582 | 0.90944
0.31742 | -33.9592 | 0.86532
0.32353 | -33.9605 | 0.82278
0.32963 | -33.9618 | 0.78113
0.33573 | -33.9632 | 0.74022
0.34184 | -33.9647 | 0.6991
0.34794 | -33.966 | 0.65804
0.35405 | -33.9672 | 0.6173
0.36015 | -33.9683 | 0.57678
0.36626 | -33.9692 | 0.53634
0.37236 | -33.97 | 0.49612
0.37846 | -33.9706 | 0.45677
0.38457 | -33.9708 | 0.41971
0.39067 | -33.9711 | 0.38232
0.39678 | -33.9714 | 0.346
0.40288 | -33.9714 | 0.31098
0.40899 | -33.9715 | 0.27636
0.41509 | -33.9716 | 0.24366
0.42119 | -33.9718 | 0.21486
0.4273 | -33.972 | 0.18374
0.4334 | -33.9722 | 0.15382
0.43951 | -33.9725 | 0.12436
0.44561 | -33.9728 | 0.096159
0.45172 | -33.9734 | 0.070063
0.45782 | -33.9739 | 0.043346
0.46392 | -33.9745 | 0.016589
0.47003 | -33.9751 | -0.0092521
0.47613 | -33.9757 | -0.034846
0.48224 | -33.9764 | -0.060151
0.48834 | -33.9769 | -0.085385
0.49445 | -33.9775 | -0.11037
0.50055 | -33.9781 | -0.13464
0.50665 | -33.9786 | -0.15928
0.51276 | -33.979 | -0.18429
0.51886 | -33.9794 | -0.20887
0.52497 | -33.9796 | -0.23284
0.53107 | -33.9799 | -0.2571
0.53717 | -33.9801 | -0.28099
0.54328 | -33.9802 | -0.30451
0.54938 | -33.9803 | -0.32813
0.55549 | -33.9803 | -0.35091
0.56159 | -33.9803 | -0.37274
0.5677 | -33.9803 | -0.39506
0.5738 | -33.9803 | -0.41725
0.5799 | -33.9803 | -0.43892
0.58601 | -33.9804 | -0.46
0.59211 | -33.9806 | -0.47945
0.59822 | -33.9808 | -0.49942
0.60432 | -33.9809 | -0.51982
0.61043 | -33.9811 | -0.53965
0.61653 | -33.9815 | -0.55885
0.62263 | -33.9818 | -0.57729
0.62874 | -33.9825 | -0.59535
0.63484 | -33.983 | -0.61392
0.64095 | -33.9836 | -0.63219
0.64705 | -33.9842 | -0.65074
0.65316 | -33.9849 | -0.66906
0.65926 | -33.9855 | -0.68705
0.66536 | -33.9864 | -0.7053
0.67147 | -33.987 | -0.72374
0.67757 | -33.9878 | -0.74182
0.68368 | -33.9886 | -0.7601
0.68978 | -33.9894 | -0.77847
0.69589 | -33.9902 | -0.79641
0.70199 | -33.9909 | -0.81416
0.70809 | -33.9917 | -0.8317
0.7142 | -33.9924 | -0.84919
0.7203 | -33.9931 | -0.86687
0.72641 | -33.9938 | -0.8843
0.73251 | -33.9944 | -0.90127
0.73862 | -33.995 | -0.91721
0.74472 | -33.9955 | -0.93343
0.75082 | -33.9961 | -0.94988
0.75693 | -33.9966 | -0.96602
0.76303 | -33.9971 | -0.98199
0.76914 | -33.9975 | -0.99644
0.77524 | -33.9979 | -1.012
0.78135 | -33.9983 | -1.0273
0.78745 | -33.9988 | -1.0426
0.79355 | -33.9992 | -1.0575
0.79966 | -33.9994 | -1.0722
0.80576 | -33.9997 | -1.0869
0.81187 | -34 | -1.1019
0.81797 | -34.0003 | -1.1167
0.82408 | -34.0006 | -1.1318
0.83018 | -34.0009 | -1.1466
0.83628 | -34.0012 | -1.1618
0.84239 | -34.0014 | -1.1775
0.84849 | -34.0017 | -1.193
0.8546 | -34.0019 | -1.2081
0.8607 | -34.0022 | -1.2233
0.86681 | -34.0024 | -1.2387
0.87291 | -34.0027 | -1.2544
0.87901 | -34.003 | -1.2707
0.88512 | -34.0032 | -1.2863
0.89122 | -34.0035 | -1.3022
0.89733 | -34.0038 | -1.3181
0.90343 | -34.0041 | -1.3336
0.90953 | -34.0044 | -1.3496
0.91564 | -34.0047 | -1.3652
0.92174 | -34.005 | -1.3809
0.92785 | -34.0053 | -1.3967
0.93395 | -34.0056 | -1.4122
0.94006 | -34.0059 | -1.4277
0.94616 | -34.0062 | -1.4429
0.95226 | -34.0066 | -1.4585
0.95837 | -34.0069 | -1.4737
0.96447 | -34.0073 | -1.4885
0.97058 | -34.0076 | -1.5036
0.97668 | -34.0079 | -1.5187
0.98279 | -34.0082 | -1.533
0.98889 | -34.0085 | -1.5475
0.99499 | -34.0088 | -1.5621
1.0011 | -34.0091 | -1.5767
1.0072 | -34.0094 | -1.5912
1.0133 | -34.0097 | -1.6053
1.0194 | -34.01 | -1.619
1.0255 | -34.0103 | -1.6333
1.0316 | -34.0105 | -1.6474
1.0377 | -34.0108 | -1.6614
1.0438 | -34.0111 | -1.6754
1.0499 | -34.0113 | -1.6886
1.056 | -34.0115 | -1.7024
1.0621 | -34.0118 | -1.7162
1.0682 | -34.0119 | -1.7295
1.0743 | -34.0122 | -1.7432
1.0805 | -34.0124 | -1.7565
1.0866 | -34.0126 | -1.7699
1.0927 | -34.0128 | -1.7832
1.0988 | -34.013 | -1.7965
1.1049 | -34.0132 | -1.81
1.111 | -34.0134 | -1.8235
1.1171 | -34.0135 | -1.8367
1.1232 | -34.0137 | -1.8499
1.1293 | -34.0139 | -1.8631
1.1354 | -34.0141 | -1.8761
1.1415 | -34.0143 | -1.8895
1.1476 | -34.0145 | -1.903
1.1537 | -34.0147 | -1.9161
1.1598 | -34.0149 | -1.929
1.1659 | -34.0151 | -1.9422
1.172 | -34.0152 | -1.9553
1.1781 | -34.0154 | -1.9684
1.1842 | -34.0156 | -1.9813
1.1903 | -34.0158 | -1.994
1.1964 | -34.016 | -2.0069
1.2025 | -34.0162 | -2.0198
1.2086 | -34.0164 | -2.0325
1.2147 | -34.0167 | -2.0453
1.2209 | -34.0169 | -2.0579
1.227 | -34.0172 | -2.0704
1.2331 | -34.0174 | -2.0832
1.2392 | -34.0176 | -2.0961
1.2453 | -34.0179 | -2.1087
1.2514 | -34.0181 | -2.121
1.2575 | -34.0184 | -2.1335
1.2636 | -34.0187 | -2.146
1.2697 | -34.019 | -2.1586
1.2758 | -34.0193 | -2.1712
1.2819 | -34.0196 | -2.1837
1.288 | -34.0199 | -2.1962
1.2941 | -34.0202 | -2.2088
1.3002 | -34.0205 | -2.2215
1.3063 | -34.0208 | -2.2339
1.3124 | -34.0211 | -2.2463
1.3185 | -34.0214 | -2.2589
1.3246 | -34.0217 | -2.2714
1.3307 | -34.022 | -2.2837
1.3368 | -34.0224 | -2.2961
1.3429 | -34.0227 | -2.3087
1.349 | -34.023 | -2.3213
1.3551 | -34.0233 | -2.3337
1.3613 | -34.0236 | -2.3462
1.3674 | -34.0239 | -2.3585
1.3735 | -34.0241 | -2.3709
1.3796 | -34.0245 | -2.3833
1.3857 | -34.0248 | -2.396
1.3918 | -34.0251 | -2.4081
1.3979 | -34.0253 | -2.4204
1.404 | -34.0256 | -2.4326
1.4101 | -34.0259 | -2.445
1.4162 | -34.0262 | -2.4574
1.4223 | -34.0265 | -2.4697
1.4284 | -34.0267 | -2.482
1.4345 | -34.027 | -2.4943
1.4406 | -34.0272 | -2.5062
1.4467 | -34.0275 | -2.5184
1.4528 | -34.0277 | -2.5306
1.4589 | -34.028 | -2.5427
1.465 | -34.0282 | -2.5547
1.4711 | -34.0284 | -2.5666
1.4772 | -34.0286 | -2.5786
1.4833 | -34.0289 | -2.5906
1.4894 | -34.0291 | -2.6025
1.4955 | -34.0293 | -2.6143
1.5016 | -34.0296 | -2.6263
1.5078 | -34.0298 | -2.6381
1.5139 | -34.0301 | -2.6499
1.52 | -34.0303 | -2.6618
1.5261 | -34.0305 | -2.6735
1.5322 | -34.0307 | -2.6855
1.5383 | -34.0309 | -2.6972
1.5444 | -34.0311 | -2.7088
1.5505 | -34.0314 | -2.7206
1.5566 | -34.0316 | -2.7323
1.5627 | -34.0318 | -2.7441
1.5688 | -34.0321 | -2.7558
1.5749 | -34.0323 | -2.7676
1.581 | -34.0325 | -2.7793
1.5871 | -34.0327 | -2.791
1.5932 | -34.0329 | -2.8028
1.5993 | -34.0331 | -2.8144
1.6054 | -34.0333 | -2.826
1.6115 | -34.0335 | -2.8377
1.6176 | -34.0337 | -2.8494
1.6237 | -34.0339 | -2.8612
1.6298 | -34.0341 | -2.8728
1.6359 | -34.0343 | -2.8844
1.642 | -34.0345 | -2.8959
1.6482 | -34.0348 | -2.9075
1.6543 | -34.035 | -2.919
1.6604 | -34.0352 | -2.9308
1.6665 | -34.0354 | -2.9424
1.6726 | -34.0357 | -2.9539
1.6787 | -34.0359 | -2.9655
1.6848 | -34.0361 | -2.9773
1.6909 | -34.0363 | -2.9889
1.697 | -34.0364 | -3.0006
1.7031 | -34.0366 | -3.0122
1.7092 | -34.0368 | -3.0238
1.7153 | -34.037 | -3.0355
1.7214 | -34.0372 | -3.0473
1.7275 | -34.0374 | -3.059
1.7336 | -34.0376 | -3.0708
1.7397 | -34.0378 | -3.0823
1.7458 | -34.0381 | -3.094
1.7519 | -34.0383 | -3.1056
1.758 | -34.0384 | -3.117
1.7641 | -34.0386 | -3.1285
1.7702 | -34.0388 | -3.1401
1.7763 | -34.039 | -3.1516
1.7824 | -34.0392 | -3.1632
1.7885 | -34.0394 | -3.1748
1.7947 | -34.0396 | -3.1864
1.8008 | -34.0398 | -3.1978
1.8069 | -34.0399 | -3.2092
1.813 | -34.0401 | -3.2208
1.8191 | -34.0403 | -3.2323
1.8252 | -34.0405 | -3.2439
1.8313 | -34.0407 | -3.2556
1.8374 | -34.0409 | -3.2672
1.8435 | -34.0411 | -3.2787
1.8496 | -34.0413 | -3.2902
1.8557 | -34.0415 | -3.3017
1.8618 | -34.0416 | -3.3133
1.8679 | -34.0418 | -3.3249
1.874 | -34.042 | -3.3365
1.8801 | -34.0422 | -3.3483
1.8862 | -34.0424 | -3.3597
1.8923 | -34.0425 | -3.3713
1.8984 | -34.0427 | -3.383
1.9045 | -34.0429 | -3.3946
1.9106 | -34.043 | -3.4063
1.9167 | -34.0432 | -3.4178
1.9228 | -34.0434 | -3.4294
1.9289 | -34.0436 | -3.4411
1.9351 | -34.0438 | -3.4526
1.9412 | -34.0439 | -3.4642
1.9473 | -34.0441 | -3.4757
1.9534 | -34.0443 | -3.4873
1.9595 | -34.0445 | -3.499
1.9656 | -34.0447 | -3.5107
1.9717 | -34.0448 | -3.5222
1.9778 | -34.045 | -3.5338
1.9839 | -34.0452 | -3.5454
1.99 | -34.0454 | -3.5571
1.9961 | -34.0456 | -3.5687
2.0022 | -34.0458 | -3.5803
2.0083 | -34.0459 | -3.5919
2.0144 | -34.0461 | -3.6033
2.0205 | -34.0463 | -3.6148
2.0266 | -34.0465 | -3.6263
2.0327 | -34.0467 | -3.6379
2.0388 | -34.0469 | -3.6494
2.0449 | -34.047 | -3.6609
2.051 | -34.0472 | -3.6724
2.0571 | -34.0474 | -3.6839
2.0632 | -34.0476 | -3.6954
2.0693 | -34.0478 | -3.707
2.0754 | -34.048 | -3.7183
2.0816 | -34.0482 | -3.7297
2.0877 | -34.0484 | -3.7412
2.0938 | -34.0486 | -3.7527
2.0999 | -34.0488 | -3.7642
2.106 | -34.049 | -3.7755
2.1121 | -34.0492 | -3.7869
2.1182 | -34.0494 | -3.7983
2.1243 | -34.0496 | -3.8097
2.1304 | -34.0498 | -3.8211
2.1365 | -34.0501 | -3.8326
2.1426 | -34.0503 | -3.8439
2.1487 | -34.0505 | -3.8553
2.1548 | -34.0507 | -3.8665
2.1609 | -34.0509 | -3.8777
2.167 | -34.0511 | -3.889
2.1731 | -34.0513 | -3.9003
2.1792 | -34.0516 | -3.9117
2.1853 | -34.0518 | -3.9231
2.1914 | -34.052 | -3.9344
2.1975 | -34.0522 | -3.9457
2.2036 | -34.0524 | -3.957
2.2097 | -34.0527 | -3.9685
2.2158 | -34.0529 | -3.9796
2.222 | -34.0531 | -3.9909
2.2281 | -34.0533 | -4.0023
2.2342 | -34.0536 | -4.0136
2.2403 | -34.0538 | -4.025
2.2464 | -34.054 | -4.0361
2.2525 | -34.0542 | -4.0474
2.2586 | -34.0544 | -4.0586
2.2647 | -34.0547 | -4.0698
2.2708 | -34.0549 | -4.0811
2.2769 | -34.0551 | -4.0923
2.283 | -34.0553 | -4.1035
2.2891 | -34.0556 | -4.1149
2.2952 | -34.0558 | -4.1261
2.3013 | -34.056 | -4.1374
2.3074 | -34.0562 | -4.1486
2.3135 | -34.0564 | -4.1599
2.3196 | -34.0566 | -4.1712
2.3257 | -34.0568 | -4.1826
2.3318 | -34.057 | -4.194
2.3379 | -34.0572 | -4.2053
2.344 | -34.0575 | -4.2166
2.3501 | -34.0576 | -4.2278
2.3562 | -34.0579 | -4.2389
2.3623 | -34.0581 | -4.2503
2.3685 | -34.0583 | -4.2618
2.3746 | -34.0585 | -4.273
2.3807 | -34.0587 | -4.2843
2.3868 | -34.0589 | -4.2955
2.3929 | -34.0591 | -4.3068
2.399 | -34.0593 | -4.3182
2.4051 | -34.0595 | -4.3293
2.4112 | -34.0597 | -4.3404
2.4173 | -34.06 | -4.3517
2.4234 | -34.0601 | -4.3631
2.4295 | -34.0603 | -4.3742
2.4356 | -34.0605 | -4.3856
2.4417 | -34.0607 | -4.3969
2.4478 | -34.0609 | -4.4081
2.4539 | -34.0611 | -4.4194
2.46 | -34.0613 | -4.4306
2.4661 | -34.0615 | -4.4418
2.4722 | -34.0617 | -4.453
2.4783 | -34.0619 | -4.4643
2.4844 | -34.0621 | -4.4755
2.4905 | -34.0622 | -4.4868
2.4966 | -34.0624 | -4.498
2.5027 | -34.0626 | -4.5093
2.5089 | -34.0628 | -4.5205
2.515 | -34.063 | -4.5318
2.5211 | -34.0632 | -4.543
2.5272 | -34.0633 | -4.5541
2.5333 | -34.0635 | -4.5652
2.5394 | -34.0637 | -4.5765
2.5455 | -34.0639 | -4.5877
2.5516 | -34.0641 | -4.5988
2.5577 | -34.0642 | -4.6099
2.5638 | -34.0644 | -4.621
2.5699 | -34.0646 | -4.6321
2.576 | -34.0648 | -4.6433
2.5821 | -34.065 | -4.6545
2.5882 | -34.0652 | -4.6656
2.5943 | -34.0654 | -4.6767
2.6004 | -34.0656 | -4.6876
2.6065 | -34.0657 | -4.6987
2.6126 | -34.0659 | -4.7097
2.6187 | -34.0661 | -4.7207
2.6248 | -34.0663 | -4.7317
2.6309 | -34.0665 | -4.7428
2.637 | -34.0667 | -4.7536
2.6431 | -34.0668 | -4.7646
2.6492 | -34.067 | -4.7756
2.6554 | -34.0672 | -4.7866
2.6615 | -34.0674 | -4.7976
2.6676 | -34.0676 | -4.8086
2.6737 | -34.0678 | -4.8193
2.6798 | -34.068 | -4.8301
2.6859 | -34.0681 | -4.841
2.692 | -34.0683 | -4.8519
2.6981 | -34.0685 | -4.8629
2.7042 | -34.0687 | -4.8737
2.7103 | -34.0689 | -4.8846
2.7164 | -34.0691 | -4.8955
2.7225 | -34.0693 | -4.9064
2.7286 | -34.0695 | -4.9172
2.7347 | -34.0697 | -4.9281
2.7408 | -34.0698 | -4.9388
2.7469 | -34.07 | -4.9497
2.753 | -34.0702 | -4.9604
2.7591 | -34.0704 | -4.9713
2.7652 | -34.0706 | -4.9821
2.7713 | -34.0708 | -4.9929
2.7774 | -34.071 | -5.0037
2.7835 | -34.0712 | -5.0144
2.7896 | -34.0714 | -5.0252
2.7958 | -34.0715 | -5.0359
2.8019 | -34.0717 | -5.0465
2.808 | -34.0719 | -5.0573
2.8141 | -34.0721 | -5.068
2.8202 | -34.0724 | -5.0788
2.8263 | -34.0726 | -5.0895
2.8324 | -34.0728 | -5.1002
2.8385 | -34.073 | -5.1109
2.8446 | -34.0732 | -5.1215
2.8507 | -34.0734 | -5.1322
2.8568 | -34.0736 | -5.1429
2.8629 | -34.0737 | -5.1536
2.869 | -34.074 | -5.1643
2.8751 | -34.0742 | -5.175
2.8812 | -34.0744 | -5.1857
2.8873 | -34.0746 | -5.1964
2.8934 | -34.0748 | -5.207
2.8995 | -34.075 | -5.2177
2.9056 | -34.0753 | -5.2284
2.9117 | -34.0755 | -5.2391
2.9178 | -34.0757 | -5.2497
2.9239 | -34.0759 | -5.2605
2.93 | -34.0761 | -5.2711
2.9361 | -34.0763 | -5.2819
2.9423 | -34.0765 | -5.2926
2.9484 | -34.0768 | -5.3033
2.9545 | -34.077 | -5.3141
2.9606 | -34.0772 | -5.3248
2.9667 | -34.0774 | -5.3357
2.9728 | -34.0776 | -5.3464
2.9789 | -34.0779 | -5.357
2.985 | -34.0781 | -5.3678
2.9911 | -34.0783 | -5.3785
2.9972 | -34.0786 | -5.3892
3.0033 | -34.0788 | -5.4001
3.0094 | -34.079 | -5.4108
3.0155 | -34.0792 | -5.4215
3.0216 | -34.0795 | -5.4323
3.0277 | -34.0797 | -5.4431
3.0338 | -34.0799 | -5.4538
3.0399 | -34.0801 | -5.4646
3.046 | -34.0804 | -5.4753
3.0521 | -34.0806 | -5.4861
3.0582 | -34.0808 | -5.4971
3.0643 | -34.081 | -5.5079
3.0704 | -34.0813 | -5.5188
3.0765 | -34.0815 | -5.5297
3.0827 | -34.0817 | -5.5406
3.0888 | -34.082 | -5.5515
3.0949 | -34.0822 | -5.5623
3.101 | -34.0824 | -5.5732
3.1071 | -34.0827 | -5.5841
3.1132 | -34.0829 | -5.5949
3.1193 | -34.0832 | -5.6057
3.1254 | -34.0834 | -5.6166
3.1315 | -34.0836 | -5.6275
3.1376 | -34.0839 | -5.6384
3.1437 | -34.0841 | -5.6493
3.1498 | -34.0843 | -5.6602
3.1559 | -34.0846 | -5.6711
3.162 | -34.0848 | -5.6821
3.1681 | -34.0851 | -5.6931
3.1742 | -34.0853 | -5.704
3.1803 | -34.0855 | -5.7148
3.1864 | -34.0858 | -5.7257
3.1925 | -34.086 | -5.7366
3.1986 | -34.0863 | -5.7476
3.2047 | -34.0865 | -5.7586
3.2108 | -34.0867 | -5.7696
3.2169 | -34.087 | -5.7806
3.223 | -34.0872 | -5.7916
3.2292 | -34.0875 | -5.8026
3.2353 | -34.0877 | -5.8136
3.2414 | -34.0879 | -5.8246
3.2475 | -34.0882 | -5.8356
3.2536 | -34.0884 | -5.8466
3.2597 | -34.0886 | -5.8575
3.2658 | -34.0889 | -5.8684
3.2719 | -34.0891 | -5.8794
3.278 | -34.0893 | -5.8903
3.2841 | -34.0895 | -5.9013
3.2902 | -34.0898 | -5.9123
3.2963 | -34.09 | -5.9232
3.3024 | -34.0902 | -5.9343
3.3085 | -34.0905 | -5.9453
3.3146 | -34.0907 | -5.9562
3.3207 | -34.091 | -5.9671
3.3268 | -34.0912 | -5.978
3.3329 | -34.0914 | -5.989
3.339 | -34.0916 | -5.9999
3.3451 | -34.0918 | -6.011
3.3512 | -34.0921 | -6.0221
3.3573 | -34.0923 | -6.033
3.3634 | -34.0925 | -6.044
3.3696 | -34.0927 | -6.055
3.3757 | -34.0929 | -6.0659
3.3818 | -34.0932 | -6.0767
3.3879 | -34.0934 | -6.0877
3.394 | -34.0936 | -6.0986
3.4001 | -34.0938 | -6.1095
3.4062 | -34.094 | -6.1205
3.4123 | -34.0943 | -6.1314
3.4184 | -34.0945 | -6.1423
3.4245 | -34.0947 | -6.1532
3.4306 | -34.0949 | -6.1642
3.4367 | -34.0951 | -6.1751
3.4428 | -34.0953 | -6.1859
3.4489 | -34.0955 | -6.1969
3.455 | -34.0957 | -6.2078
3.4611 | -34.096 | -6.2186
3.4672 | -34.0962 | -6.2295
3.4733 | -34.0964 | -6.2404
3.4794 | -34.0966 | -6.251
3.4855 | -34.0968 | -6.2619
3.4916 | -34.097 | -6.2728
3.4977 | -34.0972 | -6.2835
3.5038 | -34.0975 | -6.2944
3.5099 | -34.0977 | -6.3053
3.5161 | -34.0979 | -6.316
3.5222 | -34.0981 | -6.3268
3.5283 | -34.0983 | -6.3376
3.5344 | -34.0985 | -6.3484
3.5405 | -34.0987 | -6.3592
3.5466 | -34.0989 | -6.3699
3.5527 | -34.0991 | -6.3805
3.5588 | -34.0993 | -6.3914
3.5649 | -34.0995 | -6.4021
3.571 | -34.0997 | -6.4129
3.5771 | -34.0999 | -6.4237
3.5832 | -34.1001 | -6.4344
3.5893 | -34.1004 | -6.4451
3.5954 | -34.1006 | -6.4558
3.6015 | -34.1008 | -6.4666
3.6076 | -34.101 | -6.4773
3.6137 | -34.1012 | -6.488
3.6198 | -34.1014 | -6.4984
3.6259 | -34.1016 | -6.5091
3.632 | -34.1018 | -6.5198
3.6381 | -34.102 | -6.5305
3.6442 | -34.1022 | -6.5411
3.6503 | -34.1024 | -6.5517
3.6565 | -34.1026 | -6.5622
3.6626 | -34.1028 | -6.5729
3.6687 | -34.103 | -6.5835
3.6748 | -34.1032 | -6.5941
3.6809 | -34.1034 | -6.6046
3.687 | -34.1037 | -6.6152
3.6931 | -34.1039 | -6.6257
3.6992 | -34.1041 | -6.6364
3.7053 | -34.1043 | -6.6469
3.7114 | -34.1045 | -6.6575
3.7175 | -34.1047 | -6.668
3.7236 | -34.105 | -6.6784
3.7297 | -34.1052 | -6.689
3.7358 | -34.1054 | -6.6996
3.7419 | -34.1056 | -6.7102
3.748 | -34.1058 | -6.7208
3.7541 | -34.1061 | -6.7314
3.7602 | -34.1063 | -6.7417
3.7663 | -34.1065 | -6.7523
3.7724 | -34.1067 | -6.7627
3.7785 | -34.107 | -6.7732
3.7846 | -34.1072 | -6.7838
3.7907 | -34.1074 | -6.7943
3.7969 | -34.1077 | -6.8048
3.803 | -34.1079 | -6.8153
3.8091 | -34.1081 | -6.8259
3.8152 | -34.1084 | -6.8363
3.8213 | -34.1086 | -6.8468
3.8274 | -34.1088 | -6.8573
3.8335 | -34.1091 | -6.8677
3.8396 | -34.1093 | -6.8781
3.8457 | -34.1095 | -6.8886
3.8518 | -34.1098 | -6.8991
3.8579 | -34.11 | -6.9095
3.864 | -34.1103 | -6.92
3.8701 | -34.1105 | -6.9305
3.8762 | -34.1107 | -6.941
3.8823 | -34.111 | -6.9515
3.8884 | -34.1112 | -6.962
3.8945 | -34.1115 | -6.9724
3.9006 | -34.1118 | -6.983
3.9067 | -34.112 | -6.9935
3.9128 | -34.1123 | -7.0041
3.9189 | -34.1125 | -7.0145
3.925 | -34.1128 | -7.025
3.9311 | -34.113 | -7.0355
3.9372 | -34.1133 | -7.046
3.9434 | -34.1135 | -7.0565
3.9495 | -34.1138 | -7.067
3.9556 | -34.114 | -7.0775
3.9617 | -34.1143 | -7.088
3.9678 | -34.1145 | -7.0985
3.9739 | -34.1148 | -7.1092
3.98 | -34.1151 | -7.1198
3.9861 | -34.1153 | -7.1302
3.9922 | -34.1156 | -7.1408
3.9983 | -34.1158 | -7.1512
4.0044 | -34.1161 | -7.1617
4.0105 | -34.1164 | -7.1724
4.0166 | -34.1166 | -7.1828
4.0227 | -34.1169 | -7.1934
4.0288 | -34.1171 | -7.204
4.0349 | -34.1174 | -7.2146
4.041 | -34.1176 | -7.2252
4.0471 | -34.1179 | -7.2357
4.0532 | -34.1182 | -7.2463
4.0593 | -34.1184 | -7.2569
4.0654 | -34.1187 | -7.2675
4.0715 | -34.119 | -7.2781
4.0776 | -34.1192 | -7.2886
4.0838 | -34.1195 | -7.2991
4.0899 | -34.1198 | -7.3097
4.096 | -34.12 | -7.3203
4.1021 | -34.1203 | -7.3309
4.1082 | -34.1205 | -7.3415
4.1143 | -34.1208 | -7.3522
4.1204 | -34.1211 | -7.3628
4.1265 | -34.1213 | -7.3734
4.1326 | -34.1216 | -7.384
4.1387 | -34.1218 | -7.3945
4.1448 | -34.1221 | -7.4051
4.1509 | -34.1224 | -7.4158
4.157 | -34.1227 | -7.4264
4.1631 | -34.1229 | -7.437
4.1692 | -34.1232 | -7.4475
4.1753 | -34.1234 | -7.4581
4.1814 | -34.1237 | -7.4687
4.1875 | -34.124 | -7.4793
4.1936 | -34.1242 | -7.49
4.1997 | -34.1245 | -7.5006
4.2058 | -34.1248 | -7.5112
4.2119 | -34.125 | -7.5218
4.218 | -34.1253 | -7.5325
4.2241 | -34.1255 | -7.5431
4.2303 | -34.1258 | -7.5536
4.2364 | -34.1261 | -7.5643
4.2425 | -34.1263 | -7.5749
4.2486 | -34.1266 | -7.5856
4.2547 | -34.1268 | -7.5963
4.2608 | -34.1271 | -7.607
4.2669 | -34.1273 | -7.6176
4.273 | -34.1276 | -7.6283
4.2791 | -34.1279 | -7.6389
4.2852 | -34.1281 | -7.6495
4.2913 | -34.1284 | -7.6603
4.2974 | -34.1286 | -7.6709
4.3035 | -34.1289 | -7.6815
4.3096 | -34.1292 | -7.6922
4.3157 | -34.1294 | -7.7029
4.3218 | -34.1297 | -7.7135
4.3279 | -34.1299 | -7.7241
4.334 | -34.1302 | -7.7349
4.3401 | -34.1304 | -7.7455
4.3462 | -34.1307 | -7.7561
4.3523 | -34.131 | -7.7667
4.3584 | -34.1312 | -7.7773
4.3645 | -34.1314 | -7.7879
4.3707 | -34.1317 | -7.7986
4.3768 | -34.132 | -7.8092
4.3829 | -34.1322 | -7.8199
4.389 | -34.1324 | -7.8305
4.3951 | -34.1327 | -7.8411
4.4012 | -34.1329 | -7.8517
4.4073 | -34.1332 | -7.8624
4.4134 | -34.1334 | -7.873
4.4195 | -34.1337 | -7.8835
4.4256 | -34.134 | -7.8941
4.4317 | -34.1342 | -7.9046
4.4378 | -34.1344 | -7.9153
4.4439 | -34.1347 | -7.926
4.45 | -34.1349 | -7.9367
4.4561 | -34.1351 | -7.9473
4.4622 | -34.1354 | -7.958
4.4683 | -34.1356 | -7.9686
4.4744 | -34.1359 | -7.9792
4.4805 | -34.1361 | -7.9898
4.4866 | -34.1364 | -8.0004
4.4927 | -34.1366 | -8.011
4.4988 | -34.1368 | -8.0216
4.5049 | -34.1371 | -8.0321
4.511 | -34.1373 | -8.0427
4.5172 | -34.1376 | -8.0532
4.5233 | -34.1378 | -8.0639
4.5294 | -34.138 | -8.0745
4.5355 | -34.1383 | -8.0851
4.5416 | -34.1385 | -8.0956
4.5477 | -34.1387 | -8.1062
4.5538 | -34.139 | -8.1168
4.5599 | -34.1392 | -8.1273
4.566 | -34.1394 | -8.1378
4.5721 | -34.1397 | -8.1483
4.5782 | -34.1399 | -8.1588
4.5843 | -34.1402 | -8.1694
4.5904 | -34.1404 | -8.1799
4.5965 | -34.1407 | -8.1905
4.6026 | -34.1409 | -8.2011
4.6087 | -34.1411 | -8.2116
4.6148 | -34.1414 | -8.2222
4.6209 | -34.1416 | -8.2327
4.627 | -34.1418 | -8.2433
4.6331 | -34.1421 | -8.2538
4.6392 | -34.1423 | -8.2644
4.6453 | -34.1425 | -8.2749
4.6514 | -34.1428 | -8.2855
4.6576 | -34.143 | -8.2961
4.6637 | -34.1432 | -8.3066
4.6698 | -34.1435 | -8.3172
4.6759 | -34.1437 | -8.3277
4.682 | -34.144 | -8.3382
4.6881 | -34.1442 | -8.3488
4.6942 | -34.1444 | -8.3594
4.7003 | -34.1447 | -8.3699
4.7064 | -34.1449 | -8.3803
4.7125 | -34.1452 | -8.3907
4.7186 | -34.1454 | -8.4012
4.7247 | -34.1456 | -8.4117
4.7308 | -34.1459 | -8.4222
4.7369 | -34.1461 | -8.4327
4.743 | -34.1463 | -8.4431
4.7491 | -34.1466 | -8.4536
4.7552 | -34.1468 | -8.4641
4.7613 | -34.1471 | -8.4746
4.7674 | -34.1473 | -8.4851
4.7735 | -34.1475 | -8.4957
4.7796 | -34.1478 | -8.5061
4.7857 | -34.148 | -8.5166
4.7918 | -34.1482 | -8.5271
4.7979 | -34.1485 | -8.5376
4.8041 | -34.1487 | -8.548
4.8102 | -34.1489 | -8.5585
4.8163 | -34.1492 | -8.569
4.8224 | -34.1494 | -8.5796
4.8285 | -34.1496 | -8.59
4.8346 | -34.1499 | -8.6004
4.8407 | -34.1501 | -8.6108
4.8468 | -34.1503 | -8.6213
4.8529 | -34.1506 | -8.6317
4.859 | -34.1509 | -8.6422
4.8651 | -34.1511 | -8.6525
4.8712 | -34.1513 | -8.6629
4.8773 | -34.1516 | -8.6734
4.8834 | -34.1518 | -8.6839
4.8895 | -34.1521 | -8.6943
4.8956 | -34.1523 | -8.7048
4.9017 | -34.1525 | -8.7152
4.9078 | -34.1528 | -8.7256
4.9139 | -34.153 | -8.736
4.92 | -34.1532 | -8.7463
4.9261 | -34.1535 | -8.7568
4.9322 | -34.1537 | -8.7672
4.9383 | -34.1539 | -8.7776
4.9445 | -34.1542 | -8.7881
4.9506 | -34.1544 | -8.7985
4.9567 | -34.1547 | -8.8088
4.9628 | -34.1549 | -8.8193
4.9689 | -34.1551 | -8.8296
4.975 | -34.1554 | -8.84
4.9811 | -34.1556 | -8.8505
4.9872 | -34.1559 | -8.8608
4.9933 | -34.1561 | -8.8711
4.9994 | -34.1564 | -8.8814
5.0055 | -34.1566 | -8.8918
5.0116 | -34.1569 | -8.9022
5.0177 | -34.1571 | -8.9126
5.0238 | -34.1573 | -8.9229
5.0299 | -34.1576 | -8.9332
5.036 | -34.1578 | -8.9437
5.0421 | -34.1581 | -8.9541
5.0482 | -34.1583 | -8.9645
5.0543 | -34.1586 | -8.9749
5.0604 | -34.1588 | -8.9851
5.0665 | -34.1591 | -8.9955
5.0726 | -34.1593 | -9.0058
5.0787 | -34.1595 | -9.0163
5.0848 | -34.1598 | -9.0265
5.091 | -34.16 | -9.0369
5.0971 | -34.1603 | -9.0473
5.1032 | -34.1605 | -9.0576
5.1093 | -34.1608 | -9.0679
5.1154 | -34.161 | -9.0783
5.1215 | -34.1613 | -9.0886
5.1276 | -34.1615 | -9.0989
5.1337 | -34.1617 | -9.1091
5.1398 | -34.162 | -9.1195
5.1459 | -34.1622 | -9.1297
5.152 | -34.1625 | -9.1401
5.1581 | -34.1627 | -9.1506
5.1642 | -34.163 | -9.1609
5.1703 | -34.1633 | -9.1713
5.1764 | -34.1635 | -9.1817
5.1825 | -34.1637 | -9.1919
5.1886 | -34.164 | -9.2023
5.1947 | -34.1642 | -9.2127
5.2008 | -34.1645 | -9.223
5.2069 | -34.1648 | -9.2335
5.213 | -34.165 | -9.2439
5.2191 | -34.1653 | -9.2542
5.2252 | -34.1655 | -9.2646
5.2314 | -34.1658 | -9.275
5.2375 | -34.166 | -9.2853
5.2436 | -34.1663 | -9.2957
5.2497 | -34.1665 | -9.3061
5.2558 | -34.1668 | -9.3164
5.2619 | -34.1671 | -9.3267
5.268 | -34.1673 | -9.337
5.2741 | -34.1676 | -9.3473
5.2802 | -34.1679 | -9.3577
5.2863 | -34.1681 | -9.368
5.2924 | -34.1684 | -9.3784
5.2985 | -34.1686 | -9.3887
5.3046 | -34.1689 | -9.3991
5.3107 | -34.1692 | -9.4095
5.3168 | -34.1694 | -9.4199
5.3229 | -34.1697 | -9.4302
5.329 | -34.17 | -9.4405
5.3351 | -34.1702 | -9.4508
5.3412 | -34.1705 | -9.4612
5.3473 | -34.1707 | -9.4716
5.3534 | -34.171 | -9.4819
5.3595 | -34.1713 | -9.4924
5.3656 | -34.1716 | -9.5027
5.3717 | -34.1718 | -9.5132
5.3779 | -34.1721 | -9.5235
5.384 | -34.1724 | -9.5339
5.3901 | -34.1727 | -9.5442
5.3962 | -34.1729 | -9.5546
5.4023 | -34.1732 | -9.5649
5.4084 | -34.1735 | -9.5753
5.4145 | -34.1738 | -9.5856
5.4206 | -34.174 | -9.596
5.4267 | -34.1743 | -9.6064
5.4328 | -34.1746 | -9.6168
5.4389 | -34.1748 | -9.6271
5.445 | -34.1751 | -9.6376
5.4511 | -34.1754 | -9.6478
5.4572 | -34.1757 | -9.6581
5.4633 | -34.1759 | -9.6685
5.4694 | -34.1762 | -9.6789
5.4755 | -34.1765 | -9.6892
5.4816 | -34.1768 | -9.6996
5.4877 | -34.1771 | -9.7099
5.4938 | -34.1773 | -9.7203
5.4999 | -34.1776 | -9.7307
5.506 | -34.1779 | -9.741
5.5121 | -34.1782 | -9.7514
5.5183 | -34.1784 | -9.7619
5.5244 | -34.1787 | -9.7723
5.5305 | -34.179 | -9.7826
5.5366 | -34.1793 | -9.7931
5.5427 | -34.1796 | -9.8035
5.5488 | -34.1798 | -9.8138
5.5549 | -34.1801 | -9.8243
5.561 | -34.1804 | -9.8347
5.5671 | -34.1807 | -9.845
5.5732 | -34.181 | -9.8554
5.5793 | -34.1813 | -9.8659
5.5854 | -34.1816 | -9.8763
5.5915 | -34.1819 | -9.8867
5.5976 | -34.1822 | -9.8971
5.6037 | -34.1825 | -9.9075
5.6098 | -34.1827 | -9.9179
5.6159 | -34.183 | -9.9283
5.622 | -34.1833 | -9.9387
5.6281 | -34.1836 | -9.9491
5.6342 | -34.1839 | -9.9595
5.6403 | -34.1842 | -9.9699
5.6464 | -34.1845 | -9.9803
5.6525 | -34.1848 | -9.9907
5.6586 | -34.1851 | -10.0011
5.6648 | -34.1854 | -10.0113
5.6709 | -34.1857 | -10.0216
5.677 | -34.186 | -10.032
5.6831 | -34.1863 | -10.0425
5.6892 | -34.1866 | -10.053
5.6953 | -34.1869 | -10.0634
5.7014 | -34.1872 | -10.0738
5.7075 | -34.1875 | -10.0842
5.7136 | -34.1878 | -10.0946
5.7197 | -34.1881 | -10.105
5.7258 | -34.1884 | -10.1155
5.7319 | -34.1887 | -10.1259
5.738 | -34.189 | -10.1362
5.7441 | -34.1893 | -10.1467
5.7502 | -34.1896 | -10.157
5.7563 | -34.1899 | -10.1673
5.7624 | -34.1902 | -10.1777
5.7685 | -34.1905 | -10.1881
5.7746 | -34.1908 | -10.1985
5.7807 | -34.1911 | -10.209
5.7868 | -34.1914 | -10.2194
5.7929 | -34.1917 | -10.2298
5.799 | -34.192 | -10.2402
5.8052 | -34.1923 | -10.2506
5.8113 | -34.1926 | -10.261
5.8174 | -34.1929 | -10.2715
5.8235 | -34.1932 | -10.2818
5.8296 | -34.1935 | -10.2923
5.8357 | -34.1938 | -10.3027
5.8418 | -34.1941 | -10.3131
5.8479 | -34.1944 | -10.3236
5.854 | -34.1947 | -10.3339
5.8601 | -34.195 | -10.3443
5.8662 | -34.1953 | -10.3546
5.8723 | -34.1956 | -10.365
5.8784 | -34.1959 | -10.3755
5.8845 | -34.1962 | -10.3858
5.8906 | -34.1965 | -10.3962
5.8967 | -34.1968 | -10.4066
5.9028 | -34.1971 | -10.4169
5.9089 | -34.1974 | -10.4272
5.915 | -34.1977 | -10.4376
5.9211 | -34.198 | -10.4481
5.9272 | -34.1983 | -10.4585
5.9333 | -34.1986 | -10.4689
5.9394 | -34.1989 | -10.4793
5.9455 | -34.1992 | -10.4896
5.9517 | -34.1995 | -10.4999
5.9578 | -34.1998 | -10.5102
5.9639 | -34.2001 | -10.5206
5.97 | -34.2004 | -10.5308
5.9761 | -34.2007 | -10.5412
5.9822 | -34.201 | -10.5515
5.9883 | -34.2013 | -10.5619
5.9944 | -34.2016 | -10.5722
6.0005 | -34.2019 | -10.5826
6.0066 | -34.2022 | -10.593
6.0127 | -34.2025 | -10.6032
6.0188 | -34.2028 | -10.6137
6.0249 | -34.2031 | -10.624
6.031 | -34.2034 | -10.6345
6.0371 | -34.2037 | -10.6448
6.0432 | -34.204 | -10.6551
6.0493 | -34.2043 | -10.6653
6.0554 | -34.2046 | -10.6756
6.0615 | -34.2049 | -10.6859
6.0676 | -34.2053 | -10.6963
6.0737 | -34.2056 | -10.7066
6.0798 | -34.2059 | -10.717
6.0859 | -34.2062 | -10.7272
6.0921 | -34.2065 | -10.7375
6.0982 | -34.2068 | -10.7479
6.1043 | -34.2071 | -10.7581
6.1104 | -34.2074 | -10.7684
6.1165 | -34.2077 | -10.7785
6.1226 | -34.208 | -10.7889
6.1287 | -34.2083 | -10.7992
6.1348 | -34.2086 | -10.8095
6.1409 | -34.2089 | -10.8198
6.147 | -34.2092 | -10.8301
6.1531 | -34.2095 | -10.8402
6.1592 | -34.2098 | -10.8505
6.1653 | -34.2101 | -10.8608
6.1714 | -34.2104 | -10.871
6.1775 | -34.2107 | -10.8813
6.1836 | -34.211 | -10.8916
6.1897 | -34.2113 | -10.9018
6.1958 | -34.2117 | -10.9121
6.2019 | -34.212 | -10.9224
6.208 | -34.2123 | -10.9327
6.2141 | -34.2126 | -10.943
6.2202 | -34.2129 | -10.9532
6.2263 | -34.2132 | -10.9633
6.2325 | -34.2135 | -10.9735
6.2386 | -34.2137 | -10.9838
6.2447 | -34.2141 | -10.9941
6.2508 | -34.2144 | -11.0043
6.2569 | -34.2147 | -11.0144
6.263 | -34.215 | -11.0247
6.2691 | -34.2153 | -11.0349
6.2752 | -34.2156 | -11.0452
6.2813 | -34.2159 | -11.0555
6.2874 | -34.2162 | -11.0657
6.2935 | -34.2165 | -11.0758
6.2996 | -34.2168 | -11.0861
6.3057 | -34.2171 | -11.0963
6.3118 | -34.2174 | -11.1067
6.3179 | -34.2177 | -11.1169
6.324 | -34.218 | -11.127
6.3301 | -34.2184 | -11.1371
6.3362 | -34.2187 | -11.1472
6.3423 | -34.219 | -11.1574
6.3484 | -34.2193 | -11.1676
6.3545 | -34.2196 | -11.1777
6.3606 | -34.2199 | -11.1879
6.3667 | -34.2202 | -11.1981
6.3728 | -34.2205 | -11.2083
6.379 | -34.2208 | -11.2185
6.3851 | -34.2211 | -11.2286
6.3912 | -34.2214 | -11.2389
6.3973 | -34.2217 | -11.249
6.4034 | -34.222 | -11.2592
6.4095 | -34.2223 | -11.2694
6.4156 | -34.2226 | -11.2796
6.4217 | -34.2229 | -11.2898
6.4278 | -34.2232 | -11.3
6.4339 | -34.2235 | -11.31
6.44 | -34.2238 | -11.3201
6.4461 | -34.2241 | -11.3304
6.4522 | -34.2244 | -11.3405
6.4583 | -34.2248 | -11.3508
6.4644 | -34.2251 | -11.361
6.4705 | -34.2254 | -11.371
6.4766 | -34.2257 | -11.3812
6.4827 | -34.226 | -11.3913
6.4888 | -34.2263 | -11.4015
6.4949 | -34.2266 | -11.4117
6.501 | -34.2269 | -11.4218
6.5071 | -34.2272 | -11.4318
6.5132 | -34.2275 | -11.4419
6.5194 | -34.2278 | -11.4521
6.5255 | -34.2281 | -11.4622
6.5316 | -34.2284 | -11.4723
6.5377 | -34.2287 | -11.4824
6.5438 | -34.229 | -11.4925
6.5499 | -34.2293 | -11.5025
6.556 | -34.2296 | -11.5127
6.5621 | -34.2299 | -11.5229
6.5682 | -34.2302 | -11.533
6.5743 | -34.2305 | -11.543
6.5804 | -34.2309 | -11.5531
6.5865 | -34.2312 | -11.5632
6.5926 | -34.2315 | -11.5733
6.5987 | -34.2318 | -11.5833
6.6048 | -34.2321 | -11.5934
6.6109 | -34.2324 | -11.6033
6.617 | -34.2327 | -11.6134
6.6231 | -34.233 | -11.6235
6.6292 | -34.2333 | -11.6335
6.6353 | -34.2336 | -11.6436
6.6414 | -34.2339 | -11.6536
6.6475 | -34.2342 | -11.6635
6.6536 | -34.2345 | -11.6736
6.6597 | -34.2348 | -11.6836
6.6659 | -34.2351 | -11.6937
6.672 | -34.2355 | -11.7037
6.6781 | -34.2358 | -11.7137
6.6842 | -34.2361 | -11.7237
6.6903 | -34.2364 | -11.7338
6.6964 | -34.2367 | -11.7439
6.7025 | -34.237 | -11.7539
6.7086 | -34.2373 | -11.764
6.7147 | -34.2376 | -11.774
6.7208 | -34.2379 | -11.7841
6.7269 | -34.2382 | -11.7942
6.733 | -34.2385 | -11.8042
6.7391 | -34.2388 | -11.8142
6.7452 | -34.2391 | -11.8242
6.7513 | -34.2394 | -11.8341
6.7574 | -34.2397 | -11.8442
6.7635 | -34.24 | -11.8542
6.7696 | -34.2404 | -11.8642
6.7757 | -34.2407 | -11.8742
6.7818 | -34.241 | -11.8842
6.7879 | -34.2413 | -11.8941
6.794 | -34.2416 | -11.9041
6.8001 | -34.2419 | -11.9141
6.8063 | -34.2422 | -11.9241
6.8124 | -34.2425 | -11.9341
6.8185 | -34.2428 | -11.9441
6.8246 | -34.2432 | -11.954
6.8307 | -34.2435 | -11.9639
6.8368 | -34.2438 | -11.9739
6.8429 | -34.2441 | -11.9838
6.849 | -34.2444 | -11.9938
6.8551 | -34.2447 | -12.0037
6.8612 | -34.245 | -12.0136
6.8673 | -34.2453 | -12.0237
6.8734 | -34.2456 | -12.0336
6.8795 | -34.2459 | -12.0437
6.8856 | -34.2463 | -12.0536
6.8917 | -34.2466 | -12.0635
6.8978 | -34.2469 | -12.0735
6.9039 | -34.2472 | -12.0835
6.91 | -34.2475 | -12.0934
6.9161 | -34.2478 | -12.1033
6.9222 | -34.2481 | -12.1133
6.9283 | -34.2484 | -12.1232
6.9344 | -34.2487 | -12.1332
6.9405 | -34.2491 | -12.1432
6.9466 | -34.2494 | -12.1531
6.9528 | -34.2497 | -12.163
6.9589 | -34.25 | -12.1729
6.965 | -34.2503 | -12.1828
6.9711 | -34.2506 | -12.1927
6.9772 | -34.251 | -12.2027
6.9833 | -34.2513 | -12.2126
6.9894 | -34.2516 | -12.2226
6.9955 | -34.2519 | -12.2324
7.0016 | -34.2522 | -12.2424
7.0077 | -34.2525 | -12.2523
7.0138 | -34.2528 | -12.2622
7.0199 | -34.2532 | -12.2721
7.026 | -34.2535 | -12.2821
7.0321 | -34.2538 | -12.2919
7.0382 | -34.2541 | -12.3018
7.0443 | -34.2544 | -12.3118
7.0504 | -34.2548 | -12.3218
7.0565 | -34.2551 | -12.3317
7.0626 | -34.2554 | -12.3416
7.0687 | -34.2557 | -12.3515
7.0748 | -34.256 | -12.3614
7.0809 | -34.2563 | -12.3713
7.087 | -34.2566 | -12.3812
7.0932 | -34.257 | -12.3912
7.0993 | -34.2573 | -12.4011
7.1054 | -34.2576 | -12.4109
7.1115 | -34.2579 | -12.4207
7.1176 | -34.2582 | -12.4307
7.1237 | -34.2586 | -12.4407
7.1298 | -34.2589 | -12.4505
7.1359 | -34.2592 | -12.4604
7.142 | -34.2595 | -12.4703
7.1481 | -34.2598 | -12.4803
7.1542 | -34.2602 | -12.4902
7.1603 | -34.2605 | -12.5001
7.1664 | -34.2608 | -12.51
7.1725 | -34.2612 | -12.5199
7.1786 | -34.2615 | -12.5297
7.1847 | -34.2618 | -12.5396
7.1908 | -34.2621 | -12.5495
7.1969 | -34.2624 | -12.5595
7.203 | -34.2627 | -12.5695
7.2091 | -34.2631 | -12.5794
7.2152 | -34.2634 | -12.5892
7.2213 | -34.2637 | -12.5992
7.2274 | -34.264 | -12.609
7.2335 | -34.2644 | -12.6189
7.2397 | -34.2647 | -12.6288
7.2458 | -34.265 | -12.6386
7.2519 | -34.2653 | -12.6485
7.258 | -34.2657 | -12.6584
7.2641 | -34.266 | -12.6683
7.2702 | -34.2663 | -12.6783
7.2763 | -34.2666 | -12.6881
7.2824 | -34.267 | -12.698
7.2885 | -34.2673 | -12.708
7.2946 | -34.2676 | -12.7178
7.3007 | -34.268 | -12.7278
7.3068 | -34.2683 | -12.7378
7.3129 | -34.2686 | -12.7477
7.319 | -34.269 | -12.7576
7.3251 | -34.2693 | -12.7675
7.3312 | -34.2696 | -12.7774
7.3373 | -34.2699 | -12.7872
7.3434 | -34.2703 | -12.7971
7.3495 | -34.2706 | -12.8069
7.3556 | -34.2709 | -12.8168
7.3617 | -34.2713 | -12.8267
7.3678 | -34.2716 | -12.8365
7.3739 | -34.2719 | -12.8465
7.3801 | -34.2722 | -12.8564
7.3862 | -34.2726 | -12.8663
7.3923 | -34.2729 | -12.8761
7.3984 | -34.2733 | -12.886
7.4045 | -34.2736 | -12.8959
7.4106 | -34.2739 | -12.9057
7.4167 | -34.2743 | -12.9156
7.4228 | -34.2746 | -12.9254
7.4289 | -34.275 | -12.9353
7.435 | -34.2753 | -12.9451
7.4411 | -34.2756 | -12.9551
7.4472 | -34.276 | -12.965
7.4533 | -34.2763 | -12.9749
7.4594 | -34.2767 | -12.9849
7.4655 | -34.277 | -12.9947
7.4716 | -34.2773 | -13.0045
7.4777 | -34.2777 | -13.0145
7.4838 | -34.278 | -13.0244
7.4899 | -34.2783 | -13.0344
7.496 | -34.2787 | -13.0443
7.5021 | -34.279 | -13.0542
7.5082 | -34.2794 | -13.0641
7.5143 | -34.2797 | -13.074
7.5204 | -34.28 | -13.0838
7.5266 | -34.2804 | -13.0938
7.5327 | -34.2807 | -13.1036
7.5388 | -34.2811 | -13.1136
7.5449 | -34.2814 | -13.1234
7.551 | -34.2818 | -13.1333
7.5571 | -34.2821 | -13.1432
7.5632 | -34.2825 | -13.1531
7.5693 | -34.2828 | -13.163
7.5754 | -34.2832 | -13.1729
7.5815 | -34.2835 | -13.1829
7.5876 | -34.2838 | -13.1928
7.5937 | -34.2842 | -13.2027
7.5998 | -34.2846 | -13.2125
7.6059 | -34.2849 | -13.2223
7.612 | -34.2853 | -13.2322
7.6181 | -34.2856 | -13.2421
7.6242 | -34.286 | -13.252
7.6303 | -34.2863 | -13.2619
7.6364 | -34.2867 | -13.2718
7.6425 | -34.287 | -13.2817
7.6486 | -34.2874 | -13.2916
7.6547 | -34.2877 | -13.3014
7.6608 | -34.2881 | -13.3113
7.667 | -34.2884 | -13.3213
7.6731 | -34.2888 | -13.3312
7.6792 | -34.2892 | -13.341
7.6853 | -34.2895 | -13.3509
7.6914 | -34.2899 | -13.3608
7.6975 | -34.2902 | -13.3707
7.7036 | -34.2906 | -13.3805
7.7097 | -34.2909 | -13.3904
7.7158 | -34.2913 | -13.4004
7.7219 | -34.2917 | -13.4103
7.728 | -34.292 | -13.4202
7.7341 | -34.2924 | -13.4302
7.7402 | -34.2927 | -13.44
7.7463 | -34.2931 | -13.4499
7.7524 | -34.2935 | -13.4599
7.7585 | -34.2938 | -13.4698
7.7646 | -34.2942 | -13.4797
7.7707 | -34.2946 | -13.4897
7.7768 | -34.2949 | -13.4996
7.7829 | -34.2953 | -13.5096
7.789 | -34.2956 | -13.5195
7.7951 | -34.296 | -13.5294
7.8012 | -34.2963 | -13.5392
7.8073 | -34.2967 | -13.5491
7.8135 | -34.2971 | -13.559
7.8196 | -34.2975 | -13.5689
7.8257 | -34.2978 | -13.5788
7.8318 | -34.2982 | -13.5888
7.8379 | -34.2985 | -13.5987
7.844 | -34.2989 | -13.6086
7.8501 | -34.2993 | -13.6185
7.8562 | -34.2997 | -13.6283
7.8623 | -34.3 | -13.6382
7.8684 | -34.3004 | -13.6481
7.8745 | -34.3008 | -13.658
7.8806 | -34.3012 | -13.668
7.8867 | -34.3015 | -13.6778
7.8928 | -34.3019 | -13.6877
7.8989 | -34.3023 | -13.6976
7.905 | -34.3026 | -13.7075
7.9111 | -34.303 | -13.7174
7.9172 | -34.3034 | -13.7273
7.9233 | -34.3038 | -13.7372
7.9294 | -34.3041 | -13.7471
7.9355 | -34.3045 | -13.757
7.9416 | -34.3049 | -13.767
7.9477 | -34.3053 | -13.7768
7.9539 | -34.3056 | -13.7868
7.96 | -34.306 | -13.7968
7.9661 | -34.3064 | -13.8067
7.9722 | -34.3068 | -13.8166
7.9783 | -34.3071 | -13.8265
7.9844 | -34.3075 | -13.8364
7.9905 | -34.3079 | -13.8463
7.9966 | -34.3083 | -13.8563
8.0027 | -34.3087 | -13.8662
8.0088 | -34.309 | -13.8761
8.0149 | -34.3094 | -13.886
8.021 | -34.3098 | -13.8959
8.0271 | -34.3102 | -13.9058
8.0332 | -34.3106 | -13.9157
8.0393 | -34.3109 | -13.9256
8.0454 | -34.3113 | -13.9356
8.0515 | -34.3117 | -13.9454
8.0576 | -34.3121 | -13.9553
8.0637 | -34.3125 | -13.9653
8.0698 | -34.3128 | -13.9752
8.0759 | -34.3132 | -13.9852
8.082 | -34.3136 | -13.995
8.0881 | -34.314 | -14.005
8.0942 | -34.3143 | -14.0148
8.1004 | -34.3147 | -14.0246
8.1065 | -34.3151 | -14.0346
8.1126 | -34.3155 | -14.0445
8.1187 | -34.3159 | -14.0543
8.1248 | -34.3163 | -14.0642
8.1309 | -34.3166 | -14.0741
8.137 | -34.317 | -14.084
8.1431 | -34.3174 | -14.0938
8.1492 | -34.3178 | -14.1037
8.1553 | -34.3181 | -14.1136
8.1614 | -34.3185 | -14.1235
8.1675 | -34.3189 | -14.1334
8.1736 | -34.3193 | -14.1434
8.1797 | -34.3197 | -14.1533
8.1858 | -34.3201 | -14.1632
8.1919 | -34.3205 | -14.173
8.198 | -34.3208 | -14.1829
8.2041 | -34.3212 | -14.1927
8.2102 | -34.3216 | -14.2026
8.2163 | -34.322 | -14.2125
8.2224 | -34.3224 | -14.2224
8.2285 | -34.3228 | -14.2322
8.2346 | -34.3231 | -14.2421
8.2408 | -34.3235 | -14.2519
8.2469 | -34.3239 | -14.2618
8.253 | -34.3243 | -14.2717
8.2591 | -34.3247 | -14.2815
8.2652 | -34.3251 | -14.2913
8.2713 | -34.3255 | -14.3011
8.2774 | -34.3259 | -14.311
8.2835 | -34.3263 | -14.3209
8.2896 | -34.3266 | -14.3308
8.2957 | -34.327 | -14.3406
8.3018 | -34.3274 | -14.3504
8.3079 | -34.3278 | -14.3602
8.314 | -34.3282 | -14.37
8.3201 | -34.3286 | -14.3798
8.3262 | -34.329 | -14.3897
8.3323 | -34.3293 | -14.3995
8.3384 | -34.3297 | -14.4094
8.3445 | -34.3301 | -14.4193
8.3506 | -34.3305 | -14.4291
8.3567 | -34.3309 | -14.439
8.3628 | -34.3313 | -14.4488
8.3689 | -34.3317 | -14.4585
8.375 | -34.3321 | -14.4685
8.3812 | -34.3325 | -14.4784
8.3873 | -34.3328 | -14.4882
8.3934 | -34.3332 | -14.4981
8.3995 | -34.3336 | -14.508
8.4056 | -34.334 | -14.5177
8.4117 | -34.3344 | -14.5276
8.4178 | -34.3348 | -14.5375
8.4239 | -34.3352 | -14.5473
8.43 | -34.3356 | -14.5571
8.4361 | -34.3359 | -14.567
8.4422 | -34.3363 | -14.5767
8.4483 | -34.3367 | -14.5865
8.4544 | -34.3371 | -14.5963
8.4605 | -34.3375 | -14.6062
8.4666 | -34.3379 | -14.6161
8.4727 | -34.3383 | -14.6259
8.4788 | -34.3387 | -14.6357
8.4849 | -34.3391 | -14.6456
8.491 | -34.3395 | -14.6554
8.4971 | -34.3398 | -14.6651
8.5032 | -34.3402 | -14.6749
8.5093 | -34.3406 | -14.6846
8.5154 | -34.341 | -14.6944
8.5215 | -34.3414 | -14.7042
8.5277 | -34.3417 | -14.714
8.5338 | -34.3421 | -14.7238
8.5399 | -34.3425 | -14.7336
8.546 | -34.3429 | -14.7432
8.5521 | -34.3433 | -14.753
8.5582 | -34.3437 | -14.7628
8.5643 | -34.3441 | -14.7725
8.5704 | -34.3445 | -14.7823
8.5765 | -34.3449 | -14.7921
8.5826 | -34.3452 | -14.8018
8.5887 | -34.3456 | -14.8115
8.5948 | -34.346 | -14.8213
8.6009 | -34.3464 | -14.8311
8.607 | -34.3468 | -14.8409
8.6131 | -34.3471 | -14.8507
8.6192 | -34.3475 | -14.8603
8.6253 | -34.3479 | -14.8701
8.6314 | -34.3483 | -14.8799
8.6375 | -34.3487 | -14.8897
8.6436 | -34.3491 | -14.8994
8.6497 | -34.3494 | -14.909
8.6558 | -34.3498 | -14.9188
8.6619 | -34.3502 | -14.9286
8.6681 | -34.3506 | -14.9383
8.6742 | -34.351 | -14.948
8.6803 | -34.3513 | -14.9579
8.6864 | -34.3517 | -14.9674
8.6925 | -34.3521 | -14.9772
8.6986 | -34.3525 | -14.9869
8.7047 | -34.3529 | -14.9968
8.7108 | -34.3532 | -15.0065
8.7169 | -34.3536 | -15.0162
8.723 | -34.354 | -15.0258
8.7291 | -34.3544 | -15.0355
8.7352 | -34.3548 | -15.0453
8.7413 | -34.3552 | -15.055
8.7474 | -34.3556 | -15.0648
8.7535 | -34.356 | -15.0745
8.7596 | -34.3564 | -15.084
8.7657 | -34.3568 | -15.0937
8.7718 | -34.3571 | -15.1035
8.7779 | -34.3575 | -15.1132
8.784 | -34.3579 | -15.1229
8.7901 | -34.3583 | -15.1324
8.7962 | -34.3587 | -15.1421
8.8023 | -34.359 | -15.1518
8.8084 | -34.3594 | -15.1615
8.8146 | -34.3598 | -15.1711
8.8207 | -34.3602 | -15.1809
8.8268 | -34.3606 | -15.1903
8.8329 | -34.361 | -15.1999
8.839 | -34.3614 | -15.2096
8.8451 | -34.3617 | -15.2193
8.8512 | -34.3621 | -15.229
8.8573 | -34.3625 | -15.2388
8.8634 | -34.3629 | -15.2483
8.8695 | -34.3633 | -15.2579
8.8756 | -34.3636 | -15.2676
8.8817 | -34.364 | -15.2772
8.8878 | -34.3644 | -15.2869
8.8939 | -34.3648 | -15.2966
8.9 | -34.3652 | -15.3061
8.9061 | -34.3655 | -15.3158
8.9122 | -34.3659 | -15.3255
8.9183 | -34.3663 | -15.3351
8.9244 | -34.3667 | -15.3447
8.9305 | -34.3671 | -15.3541
8.9366 | -34.3675 | -15.3637
8.9427 | -34.3678 | -15.3734
8.9488 | -34.3682 | -15.383
8.955 | -34.3686 | -15.3927
8.9611 | -34.369 | -15.4022
8.9672 | -34.3694 | -15.4116
8.9733 | -34.3698 | -15.4212
8.9794 | -34.3702 | -15.4307
8.9855 | -34.3705 | -15.4404
8.9916 | -34.3709 | -15.45
8.9977 | -34.3713 | -15.4597
9.0038 | -34.3717 | -15.4691
9.0099 | -34.3721 | -15.4787
9.016 | -34.3725 | -15.4884
9.0221 | -34.3729 | -15.498
9.0282 | -34.3732 | -15.5076
9.0343 | -34.3736 | -15.5172
9.0404 | -34.374 | -15.5267
9.0465 | -34.3744 | -15.5364
9.0526 | -34.3748 | -15.5459
9.0587 | -34.3752 | -15.5554
9.0648 | -34.3756 | -15.565
9.0709 | -34.3759 | -15.5746
9.077 | -34.3764 | -15.584
9.0831 | -34.3767 | -15.5937
9.0892 | -34.3771 | -15.6032
9.0953 | -34.3775 | -15.6127
9.1015 | -34.3779 | -15.6223
9.1076 | -34.3783 | -15.6315
9.1137 | -34.3787 | -15.6411
9.1198 | -34.3791 | -15.6506
9.1259 | -34.3794 | -15.6602
9.132 | -34.3798 | -15.6697
9.1381 | -34.3802 | -15.6792
9.1442 | -34.3806 | -15.6885
9.1503 | -34.381 | -15.6981
9.1564 | -34.3814 | -15.7076
9.1625 | -34.3818 | -15.7171
9.1686 | -34.3822 | -15.7266
9.1747 | -34.3826 | -15.7361
9.1808 | -34.383 | -15.7455
9.1869 | -34.3834 | -15.7549
9.193 | -34.3838 | -15.7645
9.1991 | -34.3842 | -15.7741
9.2052 | -34.3845 | -15.7835
9.2113 | -34.3849 | -15.793
9.2174 | -34.3853 | -15.8023
9.2235 | -34.3857 | -15.8119
9.2296 | -34.3861 | -15.8213
9.2357 | -34.3865 | -15.8308
9.2419 | -34.3869 | -15.8403
9.248 | -34.3873 | -15.8495
9.2541 | -34.3877 | -15.8591
9.2602 | -34.3881 | -15.8685
9.2663 | -34.3885 | -15.8779
9.2724 | -34.3889 | -15.8874
9.2785 | -34.3893 | -15.8969
9.2846 | -34.3897 | -15.9061
9.2907 | -34.3901 | -15.9156
9.2968 | -34.3905 | -15.9251
9.3029 | -34.3908 | -15.9346
9.309 | -34.3912 | -15.9442
9.3151 | -34.3916 | -15.9537
9.3212 | -34.392 | -15.9629
9.3273 | -34.3924 | -15.9723
9.3334 | -34.3928 | -15.9818
9.3395 | -34.3932 | -15.9912
9.3456 | -34.3936 | -16.0007
9.3517 | -34.394 | -16.0101
9.3578 | -34.3944 | -16.0194
9.3639 | -34.3948 | -16.0288
9.37 | -34.3952 | -16.0383
9.3761 | -34.3956 | -16.0477
9.3822 | -34.396 | -16.0571
9.3884 | -34.3964 | -16.0664
9.3945 | -34.3968 | -16.0758
9.4006 | -34.3971 | -16.0853
9.4067 | -34.3975 | -16.0948
9.4128 | -34.3979 | -16.1043
9.4189 | -34.3984 | -16.1137
9.425 | -34.3988 | -16.123
9.4311 | -34.3991 | -16.1325
9.4372 | -34.3995 | -16.1419
9.4433 | -34.3999 | -16.1513
9.4494 | -34.4003 | -16.1606
9.4555 | -34.4007 | -16.17
9.4616 | -34.4011 | -16.1794
9.4677 | -34.4015 | -16.1888
9.4738 | -34.4019 | -16.1981
9.4799 | -34.4023 | -16.2075
9.486 | -34.4027 | -16.2169
9.4921 | -34.4031 | -16.2264
9.4982 | -34.4035 | -16.2357
9.5043 | -34.4039 | -16.2451
9.5104 | -34.4043 | -16.2544
9.5165 | -34.4047 | -16.2637
9.5226 | -34.4051 | -16.2731
9.5288 | -34.4055 | -16.2823
9.5349 | -34.4059 | -16.2916
9.541 | -34.4063 | -16.301
9.5471 | -34.4067 | -16.3105
9.5532 | -34.4071 | -16.3198
9.5593 | -34.4075 | -16.3291
9.5654 | -34.4079 | -16.3384
9.5715 | -34.4083 | -16.3478
9.5776 | -34.4087 | -16.3572
9.5837 | -34.4091 | -16.3667
9.5898 | -34.4095 | -16.3762
9.5959 | -34.4099 | -16.3854
9.602 | -34.4103 | -16.3948
9.6081 | -34.4107 | -16.4041
9.6142 | -34.4111 | -16.4134
9.6203 | -34.4115 | -16.4228
9.6264 | -34.4118 | -16.4322
9.6325 | -34.4123 | -16.4416
9.6386 | -34.4127 | -16.4509
9.6447 | -34.4131 | -16.4602
9.6508 | -34.4135 | -16.4696
9.6569 | -34.4139 | -16.479
9.663 | -34.4143 | -16.4883
9.6691 | -34.4147 | -16.4975
9.6753 | -34.4151 | -16.5069
9.6814 | -34.4155 | -16.5163
9.6875 | -34.4159 | -16.5256
9.6936 | -34.4163 | -16.535
9.6997 | -34.4167 | -16.5444
9.7058 | -34.4171 | -16.5535
9.7119 | -34.4175 | -16.5628
9.718 | -34.4179 | -16.5722
9.7241 | -34.4183 | -16.5814
9.7302 | -34.4187 | -16.5909
9.7363 | -34.4191 | -16.6002
9.7424 | -34.4195 | -16.6095
9.7485 | -34.4199 | -16.6188
9.7546 | -34.4203 | -16.6282
9.7607 | -34.4208 | -16.6377
9.7668 | -34.4212 | -16.647
9.7729 | -34.4216 | -16.6564
9.779 | -34.422 | -16.6655
9.7851 | -34.4224 | -16.6748
9.7912 | -34.4228 | -16.6841
9.7973 | -34.4232 | -16.6935
9.8034 | -34.4236 | -16.7028
9.8095 | -34.424 | -16.7121
9.8157 | -34.4245 | -16.7214
9.8218 | -34.4249 | -16.7307
9.8279 | -34.4253 | -16.7401
9.834 | -34.4257 | -16.7494
9.8401 | -34.4261 | -16.7587
9.8462 | -34.4265 | -16.7679
9.8523 | -34.4269 | -16.7773
9.8584 | -34.4273 | -16.7866
9.8645 | -34.4277 | -16.7958
9.8706 | -34.4281 | -16.8051
9.8767 | -34.4285 | -16.8144
9.8828 | -34.4289 | -16.8238
9.8889 | -34.4293 | -16.8331
9.895 | -34.4297 | -16.8424
9.9011 | -34.4301 | -16.8517
9.9072 | -34.4305 | -16.861
9.9133 | -34.4309 | -16.8705
9.9194 | -34.4314 | -16.8796
9.9255 | -34.4318 | -16.8889
9.9316 | -34.4322 | -16.8982
9.9377 | -34.4326 | -16.9075
9.9438 | -34.433 | -16.9168
9.9499 | -34.4334 | -16.9262
9.956 | -34.4339 | -16.9354
9.9622 | -34.4343 | -16.9447
9.9683 | -34.4347 | -16.954
9.9744 | -34.4351 | -16.9634
9.9805 | -34.4355 | -16.9726
9.9866 | -34.436 | -16.9819
9.9927 | -34.4364 | -16.9912
9.9988 | -34.4368 | -17.0005
10.0049 | -34.4372 | -17.0097
10.011 | -34.4376 | -17.0191
10.0171 | -34.438 | -17.0284
10.0232 | -34.4384 | -17.0376
10.0293 | -34.4388 | -17.0469
10.0354 | -34.4393 | -17.0561
10.0415 | -34.4397 | -17.0654
10.0476 | -34.4401 | -17.0748
10.0537 | -34.4405 | -17.0841
10.0598 | -34.4409 | -17.0935
10.0659 | -34.4413 | -17.1028
10.072 | -34.4418 | -17.1121
10.0781 | -34.4422 | -17.1215
10.0842 | -34.4426 | -17.1307
10.0903 | -34.443 | -17.14
10.0964 | -34.4434 | -17.1492
10.1026 | -34.4438 | -17.1584
10.1087 | -34.4443 | -17.1678
10.1148 | -34.4447 | -17.1771
10.1209 | -34.4451 | -17.1864
10.127 | -34.4455 | -17.1956
10.1331 | -34.4459 | -17.2049
10.1392 | -34.4463 | -17.2141
10.1453 | -34.4467 | -17.2234
10.1514 | -34.4471 | -17.2327
10.1575 | -34.4476 | -17.242
10.1636 | -34.448 | -17.2512
10.1697 | -34.4484 | -17.2605
10.1758 | -34.4488 | -17.2697
10.1819 | -34.4492 | -17.2791
10.188 | -34.4496 | -17.2884
10.1941 | -34.45 | -17.2977
10.2002 | -34.4505 | -17.3071
10.2063 | -34.4509 | -17.3163
10.2124 | -34.4513 | -17.3256
10.2185 | -34.4518 | -17.3348
10.2246 | -34.4522 | -17.3442
10.2307 | -34.4526 | -17.3535
10.2368 | -34.453 | -17.3627
10.2429 | -34.4535 | -17.3719
10.2491 | -34.4539 | -17.3812
10.2552 | -34.4543 | -17.3905
10.2613 | -34.4547 | -17.3998
10.2674 | -34.4551 | -17.409
10.2735 | -34.4556 | -17.4183
10.2796 | -34.456 | -17.4275
10.2857 | -34.4564 | -17.4368
10.2918 | -34.4568 | -17.4461
10.2979 | -34.4572 | -17.4553
10.304 | -34.4577 | -17.4645
10.3101 | -34.4581 | -17.4738
10.3162 | -34.4585 | -17.4831
10.3223 | -34.4589 | -17.4923
10.3284 | -34.4593 | -17.5015
10.3345 | -34.4597 | -17.5107
10.3406 | -34.4602 | -17.5199
10.3467 | -34.4606 | -17.5293
10.3528 | -34.461 | -17.5387
10.3589 | -34.4614 | -17.5479
10.365 | -34.4618 | -17.5571
10.3711 | -34.4623 | -17.5664
10.3772 | -34.4627 | -17.5755
10.3833 | -34.4632 | -17.5847
10.3895 | -34.4636 | -17.5939
10.3956 | -34.464 | -17.6031
10.4017 | -34.4644 | -17.6124
10.4078 | -34.4648 | -17.6216
10.4139 | -34.4653 | -17.6308
10.42 | -34.4657 | -17.64
10.4261 | -34.4661 | -17.6493
10.4322 | -34.4665 | -17.6585
10.4383 | -34.4669 | -17.6678
10.4444 | -34.4674 | -17.677
10.4505 | -34.4678 | -17.6862
10.4566 | -34.4682 | -17.6955
10.4627 | -34.4686 | -17.7047
10.4688 | -34.4691 | -17.7139
10.4749 | -34.4695 | -17.7231
10.481 | -34.4699 | -17.7323
10.4871 | -34.4703 | -17.7416
10.4932 | -34.4707 | -17.7508
10.4993 | -34.4712 | -17.7601
10.5054 | -34.4716 | -17.7694
10.5115 | -34.472 | -17.7786
10.5176 | -34.4724 | -17.7878
10.5237 | -34.4729 | -17.797
10.5298 | -34.4733 | -17.8062
10.536 | -34.4737 | -17.8155
10.5421 | -34.4741 | -17.8247
10.5482 | -34.4745 | -17.834
10.5543 | -34.475 | -17.8431
10.5604 | -34.4754 | -17.8524
10.5665 | -34.4758 | -17.8616
10.5726 | -34.4762 | -17.8708
10.5787 | -34.4767 | -17.8801
10.5848 | -34.4771 | -17.8892
10.5909 | -34.4775 | -17.8984
10.597 | -34.478 | -17.9077
10.6031 | -34.4784 | -17.917
10.6092 | -34.4788 | -17.9263
10.6153 | -34.4792 | -17.9356
10.6214 | -34.4797 | -17.9447
10.6275 | -34.4801 | -17.954
10.6336 | -34.4805 | -17.9632
10.6397 | -34.4809 | -17.9725
10.6458 | -34.4814 | -17.9816
10.6519 | -34.4818 | -17.9908
10.658 | -34.4822 | -18
10.6641 | -34.4827 | -18.0092
10.6702 | -34.4831 | -18.0184
10.6764 | -34.4835 | -18.0277
10.6825 | -34.4839 | -18.0369
10.6886 | -34.4843 | -18.0461
10.6947 | -34.4848 | -18.0553
10.7008 | -34.4852 | -18.0645
10.7069 | -34.4856 | -18.0737
10.713 | -34.4861 | -18.0829
10.7191 | -34.4865 | -18.0921
10.7252 | -34.4869 | -18.1014
10.7313 | -34.4873 | -18.1106
10.7374 | -34.4877 | -18.1199
10.7435 | -34.4882 | -18.1292
10.7496 | -34.4886 | -18.1384
10.7557 | -34.489 | -18.1475
10.7618 | -34.4894 | -18.1566
10.7679 | -34.4899 | -18.1658
10.774 | -34.4903 | -18.1749
10.7801 | -34.4907 | -18.1842
10.7862 | -34.4912 | -18.1934
10.7923 | -34.4916 | -18.2026
10.7984 | -34.492 | -18.2117
10.8045 | -34.4924 | -18.2208
10.8106 | -34.4929 | -18.2301
10.8168 | -34.4933 | -18.2394
10.8229 | -34.4937 | -18.2486
10.829 | -34.4941 | -18.2579
10.8351 | -34.4945 | -18.267
10.8412 | -34.495 | -18.2762
10.8473 | -34.4954 | -18.2853
10.8534 | -34.4958 | -18.2945
10.8595 | -34.4963 | -18.3037
10.8656 | -34.4967 | -18.3128
10.8717 | -34.4971 | -18.3221
10.8778 | -34.4976 | -18.3312
10.8839 | -34.498 | -18.3405
10.89 | -34.4984 | -18.3497
10.8961 | -34.4988 | -18.3588
10.9022 | -34.4993 | -18.368
10.9083 | -34.4997 | -18.3771
10.9144 | -34.5001 | -18.3864
10.9205 | -34.5006 | -18.3956
10.9266 | -34.501 | -18.4048
10.9327 | -34.5014 | -18.4139
10.9388 | -34.5018 | -18.4229
10.9449 | -34.5023 | -18.432
10.951 | -34.5027 | -18.4411
10.9571 | -34.5031 | -18.4504
10.9633 | -34.5035 | -18.4596
10.9694 | -34.504 | -18.4689
10.9755 | -34.5044 | -18.478
10.9816 | -34.5049 | -18.4872
10.9877 | -34.5053 | -18.4963
10.9938 | -34.5057 | -18.5055
10.9999 | -34.5062 | -18.5145
11.006 | -34.5066 | -18.5236
11.0121 | -34.507 | -18.5328
11.0182 | -34.5074 | -18.5419
11.0243 | -34.5079 | -18.5511
11.0304 | -34.5083 | -18.5603
11.0365 | -34.5087 | -18.5694
11.0426 | -34.5092 | -18.5785
11.0487 | -34.5096 | -18.5876
11.0548 | -34.51 | -18.5968
11.0609 | -34.5105 | -18.6059
11.067 | -34.5109 | -18.6152
11.0731 | -34.5113 | -18.6243
11.0792 | -34.5117 | -18.6334
11.0853 | -34.5122 | -18.6425
11.0914 | -34.5126 | -18.6516
11.0975 | -34.513 | -18.6607
11.1037 | -34.5135 | -18.6699
11.1098 | -34.5139 | -18.679
11.1159 | -34.5143 | -18.6881
11.122 | -34.5148 | -18.6973
11.1281 | -34.5152 | -18.7063
11.1342 | -34.5156 | -18.7154
11.1403 | -34.5161 | -18.7246
11.1464 | -34.5165 | -18.7337
11.1525 | -34.5169 | -18.7429
11.1586 | -34.5174 | -18.752
11.1647 | -34.5178 | -18.7611
11.1708 | -34.5182 | -18.7703
11.1769 | -34.5186 | -18.7795
11.183 | -34.5191 | -18.7884
11.1891 | -34.5195 | -18.7975
11.1952 | -34.5199 | -18.8065
11.2013 | -34.5204 | -18.8157
11.2074 | -34.5208 | -18.8249
11.2135 | -34.5213 | -18.834
11.2196 | -34.5217 | -18.8431
11.2257 | -34.5221 | -18.8523
11.2318 | -34.5226 | -18.8613
11.2379 | -34.523 | -18.8705
11.244 | -34.5235 | -18.8796
11.2502 | -34.5239 | -18.8887
11.2563 | -34.5244 | -18.8977
11.2624 | -34.5248 | -18.9069
11.2685 | -34.5252 | -18.9159
11.2746 | -34.5256 | -18.9251
11.2807 | -34.5261 | -18.9341
11.2868 | -34.5265 | -18.9433
11.2929 | -34.527 | -18.9522
11.299 | -34.5274 | -18.9612
11.3051 | -34.5278 | -18.9704
11.3112 | -34.5283 | -18.9794
11.3173 | -34.5287 | -18.9884
11.3234 | -34.5292 | -18.9974
11.3295 | -34.5296 | -19.0065
11.3356 | -34.53 | -19.0155
11.3417 | -34.5305 | -19.0247
11.3478 | -34.5309 | -19.0338
11.3539 | -34.5313 | -19.0429
11.36 | -34.5318 | -19.052
11.3661 | -34.5322 | -19.061
11.3722 | -34.5326 | -19.0701
11.3783 | -34.5331 | -19.0791
11.3844 | -34.5335 | -19.0882
11.3906 | -34.534 | -19.0973
11.3967 | -34.5344 | -19.1063
11.4028 | -34.5349 | -19.1155
11.4089 | -34.5353 | -19.1246
11.415 | -34.5357 | -19.1336
11.4211 | -34.5362 | -19.1426
11.4272 | -34.5366 | -19.1517
11.4333 | -34.5371 | -19.1607
11.4394 | -34.5375 | -19.1698
11.4455 | -34.5379 | -19.179
11.4516 | -34.5384 | -19.188
11.4577 | -34.5388 | -19.1971
11.4638 | -34.5393 | -19.206
11.4699 | -34.5397 | -19.2151
11.476 | -34.5402 | -19.2242
11.4821 | -34.5406 | -19.2332
11.4882 | -34.5411 | -19.2423
11.4943 | -34.5415 | -19.2514
11.5004 | -34.542 | -19.2603
11.5065 | -34.5424 | -19.2693
11.5126 | -34.5428 | -19.2784
11.5187 | -34.5433 | -19.2875
11.5248 | -34.5437 | -19.2965
11.5309 | -34.5441 | -19.3055
11.5371 | -34.5446 | -19.3145
11.5432 | -34.545 | -19.3236
11.5493 | -34.5455 | -19.3327
11.5554 | -34.5459 | -19.3417
11.5615 | -34.5464 | -19.3507
11.5676 | -34.5468 | -19.3598
11.5737 | -34.5473 | -19.3688
11.5798 | -34.5477 | -19.3779
11.5859 | -34.5482 | -19.3869
11.592 | -34.5486 | -19.396
11.5981 | -34.5491 | -19.4051
11.6042 | -34.5495 | -19.414
11.6103 | -34.5499 | -19.4231
11.6164 | -34.5504 | -19.4321
11.6225 | -34.5508 | -19.4412
11.6286 | -34.5513 | -19.4502
11.6347 | -34.5517 | -19.4593
11.6408 | -34.5522 | -19.4683
11.6469 | -34.5526 | -19.4773
11.653 | -34.553 | -19.4863
11.6591 | -34.5535 | -19.4954
11.6652 | -34.5539 | -19.5045
11.6713 | -34.5543 | -19.5134
11.6775 | -34.5548 | -19.5224
11.6836 | -34.5552 | -19.5314
11.6897 | -34.5557 | -19.5404
11.6958 | -34.5561 | -19.5494
11.7019 | -34.5566 | -19.5584
11.708 | -34.557 | -19.5674
11.7141 | -34.5575 | -19.5762
11.7202 | -34.558 | -19.5853
11.7263 | -34.5584 | -19.5943
11.7324 | -34.5589 | -19.6034
11.7385 | -34.5593 | -19.6124
11.7446 | -34.5598 | -19.6216
11.7507 | -34.5602 | -19.6305
11.7568 | -34.5607 | -19.6395
11.7629 | -34.5611 | -19.6485
11.769 | -34.5616 | -19.6575
11.7751 | -34.562 | -19.6665
11.7812 | -34.5625 | -19.6753
11.7873 | -34.5629 | -19.6844
11.7934 | -34.5634 | -19.6934
11.7995 | -34.5638 | -19.7024
11.8056 | -34.5643 | -19.7114
11.8117 | -34.5647 | -19.7204
11.8178 | -34.5652 | -19.7293
11.824 | -34.5656 | -19.7382
11.8301 | -34.5661 | -19.7472
11.8362 | -34.5665 | -19.7562
11.8423 | -34.567 | -19.7653
11.8484 | -34.5674 | -19.7743
11.8545 | -34.5679 | -19.7831
11.8606 | -34.5683 | -19.7921
11.8667 | -34.5688 | -19.8011
11.8728 | -34.5692 | -19.81
11.8789 | -34.5697 | -19.819
11.885 | -34.5701 | -19.8279
11.8911 | -34.5706 | -19.8368
11.8972 | -34.571 | -19.8457
11.9033 | -34.5715 | -19.8547
11.9094 | -34.5719 | -19.8636
11.9155 | -34.5724 | -19.8726
11.9216 | -34.5728 | -19.8814
11.9277 | -34.5733 | -19.8903
11.9338 | -34.5738 | -19.8993
11.9399 | -34.5742 | -19.9082
11.946 | -34.5746 | -19.9171
11.9521 | -34.5751 | -19.926
11.9582 | -34.5756 | -19.9348
11.9644 | -34.576 | -19.9438
11.9705 | -34.5765 | -19.9527
11.9766 | -34.5769 | -19.9617
11.9827 | -34.5774 | -19.9705
11.9888 | -34.5778 | -19.9793
11.9949 | -34.5783 | -19.9882
12.001 | -34.5787 | -19.9972
12.0071 | -34.5792 | -20.0061
12.0132 | -34.5796 | -20.0151
12.0193 | -34.5801 | -20.024
12.0254 | -34.5805 | -20.033
12.0315 | -34.581 | -20.0417
12.0376 | -34.5814 | -20.0507
12.0437 | -34.5819 | -20.0597
12.0498 | -34.5823 | -20.0685
12.0559 | -34.5828 | -20.0775
12.062 | -34.5832 | -20.0863
12.0681 | -34.5837 | -20.0952
12.0742 | -34.5841 | -20.1041
12.0803 | -34.5846 | -20.1131
12.0864 | -34.585 | -20.1221
12.0925 | -34.5855 | -20.131
12.0986 | -34.586 | -20.1398
12.1047 | -34.5864 | -20.1487
12.1109 | -34.5869 | -20.1576
12.117 | -34.5873 | -20.1666
12.1231 | -34.5878 | -20.1755
12.1292 | -34.5882 | -20.1845
12.1353 | -34.5887 | -20.1933
12.1414 | -34.5891 | -20.2021
12.1475 | -34.5896 | -20.2111
12.1536 | -34.59 | -20.22
12.1597 | -34.5905 | -20.2289
12.1658 | -34.591 | -20.2377
12.1719 | -34.5914 | -20.2466
12.178 | -34.5919 | -20.2555
12.1841 | -34.5923 | -20.2644
12.1902 | -34.5928 | -20.2734
12.1963 | -34.5932 | -20.2822
12.2024 | -34.5937 | -20.2909
12.2085 | -34.5942 | -20.2998
12.2146 | -34.5946 | -20.3087
12.2207 | -34.5951 | -20.3176
12.2268 | -34.5955 | -20.3265
12.2329 | -34.596 | -20.3353
12.239 | -34.5965 | -20.344
12.2451 | -34.5969 | -20.3528
12.2513 | -34.5974 | -20.3617
12.2574 | -34.5978 | -20.3706
12.2635 | -34.5983 | -20.3794
12.2696 | -34.5987 | -20.3882
12.2757 | -34.5992 | -20.397
12.2818 | -34.5996 | -20.4059
12.2879 | -34.6001 | -20.4148
12.294 | -34.6006 | -20.4236
12.3001 | -34.601 | -20.4325
12.3062 | -34.6015 | -20.4413
12.3123 | -34.6019 | -20.45
12.3184 | -34.6024 | -20.4589
12.3245 | -34.6028 | -20.4678
12.3306 | -34.6033 | -20.4767
12.3367 | -34.6037 | -20.4854
12.3428 | -34.6042 | -20.4943
12.3489 | -34.6046 | -20.503
12.355 | -34.6051 | -20.5118
12.3611 | -34.6056 | -20.5207
12.3672 | -34.606 | -20.5296
12.3733 | -34.6065 | -20.5384
12.3794 | -34.6069 | -20.5471
12.3855 | -34.6074 | -20.556
12.3916 | -34.6078 | -20.5648
12.3978 | -34.6083 | -20.5737
12.4039 | -34.6087 | -20.5826
12.41 | -34.6092 | -20.5913
12.4161 | -34.6096 | -20.5999
12.4222 | -34.6101 | -20.6087
12.4283 | -34.6106 | -20.6175
12.4344 | -34.611 | -20.6263
12.4405 | -34.6115 | -20.6352
12.4466 | -34.6119 | -20.644
12.4527 | -34.6124 | -20.6528
12.4588 | -34.6129 | -20.6616
12.4649 | -34.6133 | -20.6703
12.471 | -34.6138 | -20.6791
12.4771 | -34.6142 | -20.688
12.4832 | -34.6147 | -20.6969
12.4893 | -34.6152 | -20.7057
12.4954 | -34.6156 | -20.7145
12.5015 | -34.6161 | -20.7235
12.5076 | -34.6166 | -20.7322
12.5137 | -34.617 | -20.7409
12.5198 | -34.6175 | -20.7497
12.5259 | -34.6179 | -20.7584
12.532 | -34.6184 | -20.7673
12.5382 | -34.6189 | -20.776
12.5443 | -34.6193 | -20.7848
12.5504 | -34.6198 | -20.7936
12.5565 | -34.6203 | -20.8022
12.5626 | -34.6207 | -20.8111
12.5687 | -34.6212 | -20.8199
12.5748 | -34.6216 | -20.8286
12.5809 | -34.6221 | -20.8374
12.587 | -34.6226 | -20.8462
12.5931 | -34.623 | -20.8547
12.5992 | -34.6235 | -20.8635
12.6053 | -34.624 | -20.8723
12.6114 | -34.6244 | -20.8812
12.6175 | -34.6249 | -20.8899
12.6236 | -34.6254 | -20.8988
12.6297 | -34.6258 | -20.9074
12.6358 | -34.6263 | -20.9161
12.6419 | -34.6267 | -20.9249
12.648 | -34.6272 | -20.9337
12.6541 | -34.6276 | -20.9424
12.6602 | -34.6281 | -20.951
12.6663 | -34.6286 | -20.9598
12.6724 | -34.629 | -20.9685
12.6785 | -34.6295 | -20.9772
12.6847 | -34.6299 | -20.9861
12.6908 | -34.6304 | -20.9947
12.6969 | -34.6309 | -21.0033
12.703 | -34.6313 | -21.0121
12.7091 | -34.6318 | -21.021
12.7152 | -34.6322 | -21.0297
12.7213 | -34.6327 | -21.0385
12.7274 | -34.6332 | -21.0472
12.7335 | -34.6336 | -21.0558
12.7396 | -34.6341 | -21.0645
12.7457 | -34.6345 | -21.0733
12.7518 | -34.635 | -21.0821
12.7579 | -34.6355 | -21.0907
12.764 | -34.6359 | -21.0995
12.7701 | -34.6364 | -21.1082
12.7762 | -34.6368 | -21.1168
12.7823 | -34.6373 | -21.1255
12.7884 | -34.6378 | -21.134
12.7945 | -34.6383 | -21.1429
12.8006 | -34.6387 | -21.1515
12.8067 | -34.6392 | -21.1602
12.8128 | -34.6396 | -21.1689
12.8189 | -34.6401 | -21.1776
12.8251 | -34.6405 | -21.1864
12.8312 | -34.641 | -21.195
12.8373 | -34.6414 | -21.2037
12.8434 | -34.6419 | -21.2123
12.8495 | -34.6424 | -21.2209
12.8556 | -34.6428 | -21.2296
12.8617 | -34.6433 | -21.2383
12.8678 | -34.6438 | -21.247
12.8739 | -34.6442 | -21.2555
12.88 | -34.6447 | -21.2642
12.8861 | -34.6452 | -21.2729
12.8922 | -34.6456 | -21.2816
12.8983 | -34.6461 | -21.2903
12.9044 | -34.6466 | -21.2989
12.9105 | -34.647 | -21.3075
12.9166 | -34.6475 | -21.3163
12.9227 | -34.648 | -21.3249
12.9288 | -34.6484 | -21.3336
12.9349 | -34.6489 | -21.3421
12.941 | -34.6494 | -21.3507
12.9471 | -34.6498 | -21.3594
12.9532 | -34.6503 | -21.3681
12.9593 | -34.6508 | -21.3768
12.9654 | -34.6512 | -21.3854
12.9716 | -34.6517 | -21.3942
12.9777 | -34.6522 | -21.4028
12.9838 | -34.6527 | -21.4116
12.9899 | -34.6531 | -21.4204
12.996 | -34.6536 | -21.4291
13.0021 | -34.6541 | -21.4378
13.0082 | -34.6545 | -21.4464
13.0143 | -34.655 | -21.4548
13.0204 | -34.6554 | -21.4634
13.0265 | -34.6559 | -21.472
13.0326 | -34.6564 | -21.4807
13.0387 | -34.6568 | -21.4894
13.0448 | -34.6573 | -21.4981
13.0509 | -34.6577 | -21.5065
13.057 | -34.6582 | -21.5149
13.0631 | -34.6587 | -21.5233
13.0692 | -34.6591 | -21.5317
13.0753 | -34.6596 | -21.5402
13.0814 | -34.6601 | -21.5488
13.0875 | -34.6605 | -21.5572
13.0936 | -34.661 | -21.5658
13.0997 | -34.6615 | -21.5744
13.1058 | -34.6619 | -21.583
13.112 | -34.6624 | -21.5914
13.1181 | -34.6628 | -21.5996
13.1242 | -34.6633 | -21.6083
13.1303 | -34.6637 | -21.617
13.1364 | -34.6642 | -21.6256
13.1425 | -34.6647 | -21.6341
13.1486 | -34.6651 | -21.6427
13.1547 | -34.6656 | -21.6513
13.1608 | -34.6661 | -21.66
13.1669 | -34.6666 | -21.6687
13.173 | -34.667 | -21.6773
13.1791 | -34.6675 | -21.686
13.1852 | -34.668 | -21.6945
13.1913 | -34.6684 | -21.7029
13.1974 | -34.6689 | -21.7116
13.2035 | -34.6694 | -21.7203
13.2096 | -34.6698 | -21.7289
13.2157 | -34.6703 | -21.7375
13.2218 | -34.6708 | -21.746
13.2279 | -34.6712 | -21.7543
13.234 | -34.6717 | -21.7629
13.2401 | -34.6722 | -21.7715
13.2462 | -34.6726 | -21.7801
13.2524 | -34.6731 | -21.7887
13.2585 | -34.6736 | -21.7972
13.2646 | -34.674 | -21.8059
13.2707 | -34.6745 | -21.8145
13.2768 | -34.6749 | -21.8232
13.2829 | -34.6754 | -21.8318
13.289 | -34.6759 | -21.8405
13.2951 | -34.6764 | -21.8487
13.3012 | -34.6768 | -21.8574
13.3073 | -34.6773 | -21.866
13.3134 | -34.6777 | -21.8744
13.3195 | -34.6781 | -21.8832
13.3256 | -34.6786 | -21.8918
13.3317 | -34.679 | -21.9002
13.3378 | -34.6794 | -21.909
13.3439 | -34.6799 | -21.9173
13.35 | -34.6803 | -21.9261
13.3561 | -34.6808 | -21.9347
13.3622 | -34.6813 | -21.943
13.3683 | -34.6818 | -21.952
13.3744 | -34.6822 | -21.9606
13.3805 | -34.6826 | -21.969
13.3866 | -34.6831 | -21.9774
13.3927 | -34.6835 | -21.9864
13.3989 | -34.684 | -21.9948
13.405 | -34.6845 | -22.0034
13.4111 | -34.6849 | -22.0121
13.4172 | -34.6854 | -22.0207
13.4233 | -34.6859 | -22.0298
13.4294 | -34.6864 | -22.0385
13.4355 | -34.6869 | -22.0473
13.4416 | -34.6874 | -22.0557
13.4477 | -34.6878 | -22.064
13.4538 | -34.6883 | -22.0724
13.4599 | -34.6887 | -22.081
13.466 | -34.6892 | -22.0895
13.4721 | -34.6895 | -22.0979
13.4782 | -34.69 | -22.1066
13.4843 | -34.6905 | -22.1151
13.4904 | -34.6909 | -22.1235
13.4965 | -34.6914 | -22.1321
13.5026 | -34.6919 | -22.1409
13.5087 | -34.6924 | -22.1495
13.5148 | -34.6929 | -22.1581
13.5209 | -34.6934 | -22.1667
13.527 | -34.6938 | -22.1754
13.5331 | -34.6943 | -22.1837
13.5393 | -34.6947 | -22.1921
13.5454 | -34.6952 | -22.2005
13.5515 | -34.6956 | -22.2092
13.5576 | -34.6961 | -22.2178
13.5637 | -34.6966 | -22.2259
13.5698 | -34.6971 | -22.2345
13.5759 | -34.6976 | -22.243
13.582 | -34.6981 | -22.2514
13.5881 | -34.6986 | -22.2598
13.5942 | -34.699 | -22.2682
13.6003 | -34.6995 | -22.2767
13.6064 | -34.7 | -22.2853
13.6125 | -34.7005 | -22.2938
13.6186 | -34.701 | -22.3027
13.6247 | -34.7014 | -22.3111
13.6308 | -34.7019 | -22.3199
13.6369 | -34.7024 | -22.3288
13.643 | -34.7029 | -22.3371
13.6491 | -34.7034 | -22.3458
13.6552 | -34.7038 | -22.3544
13.6613 | -34.7043 | -22.363
13.6674 | -34.7048 | -22.3716
13.6735 | -34.7053 | -22.3802
13.6796 | -34.7057 | -22.3888
13.6858 | -34.7062 | -22.3973
13.6919 | -34.7067 | -22.4056
13.698 | -34.7071 | -22.4141
13.7041 | -34.7076 | -22.4228
13.7102 | -34.7081 | -22.4314
13.7163 | -34.7085 | -22.4398
13.7224 | -34.7091 | -22.4484
13.7285 | -34.7095 | -22.4569
13.7346 | -34.71 | -22.4654
13.7407 | -34.7105 | -22.4738
13.7468 | -34.7109 | -22.4824
13.7529 | -34.7114 | -22.4911
13.759 | -34.7119 | -22.4993
13.7651 | -34.7124 | -22.5078
13.7712 | -34.7128 | -22.5163
13.7773 | -34.7133 | -22.525
13.7834 | -34.7138 | -22.5335
13.7895 | -34.7142 | -22.5419
13.7956 | -34.7147 | -22.5508
13.8017 | -34.7152 | -22.5595
13.8078 | -34.7156 | -22.5679
13.8139 | -34.7161 | -22.5766
13.82 | -34.7166 | -22.5852
13.8262 | -34.7171 | -22.5938
13.8323 | -34.7176 | -22.6021
13.8384 | -34.718 | -22.6106
13.8445 | -34.7185 | -22.6191
13.8506 | -34.719 | -22.6276
13.8567 | -34.7195 | -22.636
13.8628 | -34.7199 | -22.6444
13.8689 | -34.7204 | -22.6528
13.875 | -34.7208 | -22.6612
13.8811 | -34.7213 | -22.6698
13.8872 | -34.7217 | -22.6784
13.8933 | -34.7222 | -22.6869
13.8994 | -34.7227 | -22.6952
13.9055 | -34.7232 | -22.7036
13.9116 | -34.7236 | -22.7122
13.9177 | -34.7241 | -22.7207
13.9238 | -34.7245 | -22.7291
13.9299 | -34.725 | -22.7376
13.936 | -34.7255 | -22.7461
13.9421 | -34.726 | -22.7546
13.9482 | -34.7265 | -22.7632
13.9543 | -34.7269 | -22.7719
13.9604 | -34.7274 | -22.7804
13.9665 | -34.7279 | -22.7889
13.9727 | -34.7284 | -22.7974
13.9788 | -34.7289 | -22.8061
13.9849 | -34.7293 | -22.8147
13.991 | -34.7298 | -22.8232
13.9971 | -34.7303 | -22.8317
14.0032 | -34.7308 | -22.8402
14.0093 | -34.7313 | -22.8487
14.0154 | -34.7318 | -22.8573
14.0215 | -34.7322 | -22.8656
14.0276 | -34.7327 | -22.8742
14.0337 | -34.7332 | -22.8825
14.0398 | -34.7337 | -22.891
14.0459 | -34.7342 | -22.8996
14.052 | -34.7347 | -22.908
14.0581 | -34.7351 | -22.9166
14.0642 | -34.7356 | -22.9252
14.0703 | -34.7361 | -22.9335
14.0764 | -34.7366 | -22.942
14.0825 | -34.737 | -22.9504
14.0886 | -34.7375 | -22.9588
14.0947 | -34.738 | -22.9673
14.1008 | -34.7384 | -22.9758
14.1069 | -34.7389 | -22.9842
14.1131 | -34.7394 | -22.9927
14.1192 | -34.7399 | -23.0012
14.1253 | -34.7404 | -23.0098
14.1314 | -34.7409 | -23.0181
14.1375 | -34.7414 | -23.0265
14.1436 | -34.7418 | -23.0349
14.1497 | -34.7423 | -23.0434
14.1558 | -34.7427 | -23.052
14.1619 | -34.7432 | -23.0606
14.168 | -34.7437 | -23.0691
14.1741 | -34.7442 | -23.0776
14.1802 | -34.7447 | -23.086
14.1863 | -34.7451 | -23.0944
14.1924 | -34.7456 | -23.103
14.1985 | -34.7461 | -23.1116
14.2046 | -34.7466 | -23.1201
14.2107 | -34.7471 | -23.1283
14.2168 | -34.7476 | -23.137
14.2229 | -34.7481 | -23.1457
14.229 | -34.7485 | -23.1542
14.2351 | -34.749 | -23.1626
14.2412 | -34.7495 | -23.1711
14.2473 | -34.75 | -23.1796
14.2534 | -34.7504 | -23.1881
14.2596 | -34.7509 | -23.1966
14.2657 | -34.7514 | -23.205
14.2718 | -34.7518 | -23.2134
14.2779 | -34.7523 | -23.2219
14.284 | -34.7528 | -23.2301
14.2901 | -34.7533 | -23.2384
14.2962 | -34.7538 | -23.2469
14.3023 | -34.7543 | -23.2554
14.3084 | -34.7547 | -23.264
14.3145 | -34.7552 | -23.2724
14.3206 | -34.7557 | -23.2808
14.3267 | -34.7562 | -23.2893
14.3328 | -34.7567 | -23.2978
14.3389 | -34.7571 | -23.3063
14.345 | -34.7576 | -23.3147
14.3511 | -34.7581 | -23.323
14.3572 | -34.7586 | -23.3316
14.3633 | -34.7591 | -23.3401
14.3694 | -34.7596 | -23.3486
14.3755 | -34.76 | -23.357
14.3816 | -34.7605 | -23.3654
14.3877 | -34.761 | -23.3738
14.3938 | -34.7615 | -23.3824
14.4 | -34.762 | -23.3908
14.4061 | -34.7625 | -23.3993
14.4122 | -34.763 | -23.4078
14.4183 | -34.7634 | -23.4163
14.4244 | -34.7639 | -23.4247
14.4305 | -34.7644 | -23.433
14.4366 | -34.7649 | -23.4414
14.4427 | -34.7654 | -23.4499
14.4488 | -34.7659 | -23.4584
14.4549 | -34.7664 | -23.4668
14.461 | -34.7669 | -23.4752
14.4671 | -34.7674 | -23.4837
14.4732 | -34.7679 | -23.4921
14.4793 | -34.7683 | -23.5006
14.4854 | -34.7688 | -23.5091
14.4915 | -34.7693 | -23.5175
14.4976 | -34.7698 | -23.5259
14.5037 | -34.7703 | -23.5344
14.5098 | -34.7708 | -23.5429
14.5159 | -34.7713 | -23.5514
14.522 | -34.7717 | -23.5599
14.5281 | -34.7722 | -23.5682
14.5342 | -34.7727 | -23.5766
14.5403 | -34.7732 | -23.5851
14.5465 | -34.7737 | -23.5935
14.5526 | -34.7742 | -23.6021
14.5587 | -34.7746 | -23.6106
14.5648 | -34.7751 | -23.6189
14.5709 | -34.7756 | -23.6274
14.577 | -34.7761 | -23.6359
14.5831 | -34.7766 | -23.6443
14.5892 | -34.7771 | -23.6526
14.5953 | -34.7776 | -23.6609
14.6014 | -34.7781 | -23.6695
14.6075 | -34.7786 | -23.6779
14.6136 | -34.779 | -23.6865
14.6197 | -34.7795 | -23.6949
14.6258 | -34.78 | -23.7033
14.6319 | -34.7805 | -23.7116
14.638 | -34.781 | -23.7201
14.6441 | -34.7815 | -23.7287
14.6502 | -34.782 | -23.7371
14.6563 | -34.7825 | -23.7456
14.6624 | -34.7829 | -23.754
14.6685 | -34.7835 | -23.7622
14.6746 | -34.7839 | -23.7707
14.6807 | -34.7844 | -23.7791
14.6869 | -34.7849 | -23.7876
14.693 | -34.7854 | -23.7959
14.6991 | -34.7859 | -23.8044
14.7052 | -34.7864 | -23.8127
14.7113 | -34.7869 | -23.821
14.7174 | -34.7874 | -23.8293
14.7235 | -34.7879 | -23.8378
14.7296 | -34.7883 | -23.8463
14.7357 | -34.7889 | -23.8545
14.7418 | -34.7893 | -23.8629
14.7479 | -34.7898 | -23.8712
14.754 | -34.7903 | -23.8797
14.7601 | -34.7908 | -23.8883
14.7662 | -34.7913 | -23.8967
14.7723 | -34.7918 | -23.9049
14.7784 | -34.7923 | -23.9134
14.7845 | -34.7927 | -23.9218
14.7906 | -34.7932 | -23.9302
14.7967 | -34.7937 | -23.9387
14.8028 | -34.7942 | -23.9471
14.8089 | -34.7947 | -23.9553
14.815 | -34.7952 | -23.9638
14.8211 | -34.7957 | -23.9721
14.8272 | -34.7962 | -23.9805
14.8334 | -34.7967 | -23.9888
14.8395 | -34.7972 | -23.9971
14.8456 | -34.7977 | -24.0054
14.8517 | -34.7982 | -24.0139
14.8578 | -34.7987 | -24.0223
14.8639 | -34.7992 | -24.0306
14.87 | -34.7997 | -24.0389
14.8761 | -34.8002 | -24.0471
14.8822 | -34.8007 | -24.0555
14.8883 | -34.8012 | -24.0639
14.8944 | -34.8017 | -24.0722
14.9005 | -34.8021 | -24.0805
14.9066 | -34.8026 | -24.0888
14.9127 | -34.8031 | -24.0971
14.9188 | -34.8036 | -24.1055
14.9249 | -34.8041 | -24.1138
14.931 | -34.8046 | -24.1222
14.9371 | -34.8051 | -24.1305
14.9432 | -34.8056 | -24.1388
14.9493 | -34.8061 | -24.1471
14.9554 | -34.8066 | -24.1554
14.9615 | -34.8071 | -24.1637
14.9676 | -34.8076 | -24.1722
14.9738 | -34.8081 | -24.1805
14.9799 | -34.8086 | -24.1888
14.986 | -34.8091 | -24.1968
14.9921 | -34.8096 | -24.2052
14.9982 | -34.8101 | -24.2135
15.0043 | -34.8105 | -24.2219
15.0104 | -34.811 | -24.2303
15.0165 | -34.8115 | -24.2385
15.0226 | -34.812 | -24.2467
15.0287 | -34.8125 | -24.255
15.0348 | -34.813 | -24.2633
15.0409 | -34.8135 | -24.2716
15.047 | -34.814 | -24.28
15.0531 | -34.8145 | -24.288
15.0592 | -34.815 | -24.2964
15.0653 | -34.8155 | -24.3046
15.0714 | -34.816 | -24.313
15.0775 | -34.8165 | -24.3213
15.0836 | -34.817 | -24.3295
15.0897 | -34.8175 | -24.3377
15.0958 | -34.818 | -24.346
15.1019 | -34.8185 | -24.3543
15.108 | -34.819 | -24.3624
15.1141 | -34.8195 | -24.3708
15.1203 | -34.82 | -24.379
15.1264 | -34.8205 | -24.387
15.1325 | -34.821 | -24.3954
15.1386 | -34.8215 | -24.4036
15.1447 | -34.822 | -24.4119
15.1508 | -34.8225 | -24.4201
15.1569 | -34.823 | -24.4284
15.163 | -34.8235 | -24.4364
15.1691 | -34.824 | -24.4448
15.1752 | -34.8245 | -24.4531
15.1813 | -34.825 | -24.4614
15.1874 | -34.8255 | -24.4697
15.1935 | -34.826 | -24.4776
15.1996 | -34.8265 | -24.4859
15.2057 | -34.827 | -24.4942
15.2118 | -34.8275 | -24.5025
15.2179 | -34.828 | -24.5107
15.224 | -34.8285 | -24.519
15.2301 | -34.829 | -24.5271
15.2362 | -34.8295 | -24.5354
15.2423 | -34.83 | -24.5437
15.2484 | -34.8305 | -24.5519
15.2545 | -34.831 | -24.5602
15.2607 | -34.8315 | -24.5684
15.2668 | -34.832 | -24.5765
15.2729 | -34.8325 | -24.5847
15.279 | -34.833 | -24.593
15.2851 | -34.8335 | -24.6012
15.2912 | -34.834 | -24.6096
15.2973 | -34.8345 | -24.6178
15.3034 | -34.835 | -24.6258
15.3095 | -34.8355 | -24.6339
15.3156 | -34.836 | -24.6422
15.3217 | -34.8365 | -24.6503
15.3278 | -34.837 | -24.6585
15.3339 | -34.8375 | -24.6665
15.34 | -34.838 | -24.6746
15.3461 | -34.8385 | -24.6829
15.3522 | -34.839 | -24.6911
15.3583 | -34.8395 | -24.6993
15.3644 | -34.84 | -24.7074
15.3705 | -34.8405 | -24.7152
15.3766 | -34.841 | -24.7234
15.3827 | -34.8415 | -24.7316
15.3888 | -34.842 | -24.7397
15.3949 | -34.8425 | -24.7478
15.401 | -34.843 | -24.756
15.4072 | -34.8435 | -24.7639
15.4133 | -34.844 | -24.7721
15.4194 | -34.8445 | -24.7803
15.4255 | -34.845 | -24.7884
15.4316 | -34.8455 | -24.7966
15.4377 | -34.846 | -24.8048
15.4438 | -34.8465 | -24.8127
15.4499 | -34.847 | -24.8208
15.456 | -34.8475 | -24.829
15.4621 | -34.848 | -24.837
15.4682 | -34.8485 | -24.8451
15.4743 | -34.849 | -24.8531
15.4804 | -34.8495 | -24.8612
15.4865 | -34.85 | -24.8694
15.4926 | -34.8505 | -24.8776
15.4987 | -34.8511 | -24.8856
15.5048 | -34.8516 | -24.8937
15.5109 | -34.8521 | -24.9016
15.517 | -34.8526 | -24.9097
15.5231 | -34.8531 | -24.9178
15.5292 | -34.8536 | -24.926
15.5353 | -34.8541 | -24.9342
15.5414 | -34.8546 | -24.9423
15.5476 | -34.8551 | -24.95
15.5537 | -34.8556 | -24.9581
15.5598 | -34.8561 | -24.9663
15.5659 | -34.8566 | -24.9744
15.572 | -34.8571 | -24.9825
15.5781 | -34.8576 | -24.9906
15.5842 | -34.8581 | -24.9986
15.5903 | -34.8586 | -25.0065
15.5964 | -34.8591 | -25.0145
15.6025 | -34.8596 | -25.0226
15.6086 | -34.8601 | -25.0306
15.6147 | -34.8606 | -25.0387
15.6208 | -34.8611 | -25.0464
15.6269 | -34.8616 | -25.0545
15.633 | -34.8621 | -25.0626
15.6391 | -34.8626 | -25.0706
15.6452 | -34.8631 | -25.0786
15.6513 | -34.8636 | -25.0864
15.6574 | -34.8641 | -25.0945
15.6635 | -34.8646 | -25.1025
15.6696 | -34.8651 | -25.1105
15.6757 | -34.8656 | -25.1185
15.6818 | -34.8661 | -25.1265
15.688 | -34.8666 | -25.1344
15.6941 | -34.8671 | -25.1425
15.7002 | -34.8676 | -25.1504
15.7063 | -34.8681 | -25.1584
15.7124 | -34.8686 | -25.1664
15.7185 | -34.8691 | -25.1745
15.7246 | -34.8696 | -25.1823
15.7307 | -34.8701 | -25.1902
15.7368 | -34.8706 | -25.1983
15.7429 | -34.8711 | -25.2062
15.749 | -34.8716 | -25.2141
15.7551 | -34.8721 | -25.2223
15.7612 | -34.8727 | -25.23
15.7673 | -34.8732 | -25.238
15.7734 | -34.8737 | -25.2458
15.7795 | -34.8741 | -25.2537
15.7856 | -34.8746 | -25.2618
15.7917 | -34.8752 | -25.2695
15.7978 | -34.8757 | -25.2775
15.8039 | -34.8761 | -25.2855
15.81 | -34.8766 | -25.2935
15.8161 | -34.8771 | -25.3015
15.8222 | -34.8777 | -25.3096
15.8283 | -34.8782 | -25.3172
15.8345 | -34.8787 | -25.3252
15.8406 | -34.8792 | -25.3332
15.8467 | -34.8797 | -25.341
15.8528 | -34.8802 | -25.349
15.8589 | -34.8807 | -25.3569
15.865 | -34.8812 | -25.3645
15.8711 | -34.8817 | -25.3725
15.8772 | -34.8822 | -25.3804
15.8833 | -34.8827 | -25.3882
15.8894 | -34.8832 | -25.3962
15.8955 | -34.8837 | -25.4042
15.9016 | -34.8841 | -25.4117
15.9077 | -34.8847 | -25.4197
15.9138 | -34.8851 | -25.4278
15.9199 | -34.8856 | -25.4355
15.926 | -34.8861 | -25.4434
15.9321 | -34.8867 | -25.4509
15.9382 | -34.8871 | -25.4588
15.9443 | -34.8877 | -25.4667
15.9504 | -34.8882 | -25.4745
15.9565 | -34.8886 | -25.4823
15.9626 | -34.8892 | -25.4902
15.9687 | -34.8896 | -25.4978
15.9749 | -34.8901 | -25.5055
15.981 | -34.8906 | -25.5134
15.9871 | -34.8911 | -25.5215
15.9932 | -34.8916 | -25.5292
15.9993 | -34.8921 | -25.537
16.0054 | -34.8926 | -25.5447
16.0115 | -34.8931 | -25.5524
16.0176 | -34.8936 | -25.5603
16.0237 | -34.8941 | -25.5682
16.0298 | -34.8946 | -25.5762
16.0359 | -34.8952 | -25.5841
16.042 | -34.8957 | -25.5916
16.0481 | -34.8962 | -25.5994
16.0542 | -34.8967 | -25.6072
16.0603 | -34.8972 | -25.6151
16.0664 | -34.8976 | -25.6232
16.0725 | -34.8982 | -25.6308
16.0786 | -34.8986 | -25.6386
16.0847 | -34.8992 | -25.6464
16.0908 | -34.8996 | -25.6542
16.0969 | -34.9001 | -25.6621
16.103 | -34.9006 | -25.6698
16.1091 | -34.9011 | -25.6772
16.1152 | -34.9016 | -25.6851
16.1214 | -34.9021 | -25.6929
16.1275 | -34.9026 | -25.7007
16.1336 | -34.9031 | -25.7086
16.1397 | -34.9036 | -25.7163
16.1458 | -34.9041 | -25.7238
16.1519 | -34.9046 | -25.7316
16.158 | -34.9051 | -25.7394
16.1641 | -34.9056 | -25.7471
16.1702 | -34.9061 | -25.7548
16.1763 | -34.9066 | -25.7626
16.1824 | -34.9071 | -25.7701
16.1885 | -34.9076 | -25.778
16.1946 | -34.9081 | -25.7858
16.2007 | -34.9086 | -25.7936
16.2068 | -34.9091 | -25.8013
16.2129 | -34.9096 | -25.8091
16.219 | -34.9101 | -25.8166
16.2251 | -34.9106 | -25.8244
16.2312 | -34.9111 | -25.832
16.2373 | -34.9116 | -25.8396
16.2434 | -34.9121 | -25.8475
16.2495 | -34.9126 | -25.855
16.2556 | -34.9131 | -25.8627
16.2618 | -34.9136 | -25.8703
16.2679 | -34.914 | -25.8781
16.274 | -34.9145 | -25.8857
16.2801 | -34.915 | -25.8936
16.2862 | -34.9155 | -25.9009
16.2923 | -34.916 | -25.9087
16.2984 | -34.9165 | -25.9164
16.3045 | -34.917 | -25.9241
16.3106 | -34.9175 | -25.9318
16.3167 | -34.918 | -25.9396
16.3228 | -34.9185 | -25.9468
16.3289 | -34.919 | -25.9547
16.335 | -34.9195 | -25.9623
16.3411 | -34.92 | -25.97
16.3472 | -34.9205 | -25.9777
16.3533 | -34.921 | -25.9854
16.3594 | -34.9215 | -25.9929
16.3655 | -34.922 | -26.0005
16.3716 | -34.9225 | -26.0082
16.3777 | -34.923 | -26.0158
16.3838 | -34.9235 | -26.0235
16.3899 | -34.924 | -26.0307
16.396 | -34.9245 | -26.0384
16.4021 | -34.925 | -26.0462
16.4083 | -34.9255 | -26.054
16.4144 | -34.926 | -26.0616
16.4205 | -34.9265 | -26.0693
16.4266 | -34.927 | -26.0768
16.4327 | -34.9275 | -26.0844
16.4388 | -34.928 | -26.0922
16.4449 | -34.9285 | -26.0996
16.451 | -34.929 | -26.1073
16.4571 | -34.9295 | -26.115
16.4632 | -34.93 | -26.1223
16.4693 | -34.9305 | -26.1299
16.4754 | -34.931 | -26.1377
16.4815 | -34.9314 | -26.1452
16.4876 | -34.932 | -26.1528
16.4937 | -34.9325 | -26.1604
16.4998 | -34.933 | -26.1678
16.5059 | -34.9335 | -26.1754
16.512 | -34.934 | -26.1831
16.5181 | -34.9345 | -26.1906
16.5242 | -34.935 | -26.1985
16.5303 | -34.9355 | -26.2058
16.5364 | -34.936 | -26.2137
16.5425 | -34.9365 | -26.2212
16.5487 | -34.937 | -26.2287
16.5548 | -34.9375 | -26.2363
16.5609 | -34.938 | -26.2438
16.567 | -34.9385 | -26.2512
16.5731 | -34.939 | -26.2588
16.5792 | -34.9395 | -26.2665
16.5853 | -34.94 | -26.274
16.5914 | -34.9405 | -26.2817
16.5975 | -34.941 | -26.2891
16.6036 | -34.9415 | -26.2964
16.6097 | -34.942 | -26.3041
16.6158 | -34.9425 | -26.3116
16.6219 | -34.943 | -26.3191
16.628 | -34.9435 | -26.3267
16.6341 | -34.944 | -26.3341
16.6402 | -34.9445 | -26.3417
16.6463 | -34.945 | -26.3493
16.6524 | -34.9455 | -26.3569
16.6585 | -34.946 | -26.3645
16.6646 | -34.9465 | -26.3721
16.6707 | -34.947 | -26.3794
16.6768 | -34.9475 | -26.387
16.6829 | -34.9479 | -26.3944
16.689 | -34.9484 | -26.4019
16.6952 | -34.9489 | -26.4095
16.7013 | -34.9494 | -26.4169
16.7074 | -34.95 | -26.4244
16.7135 | -34.9504 | -26.4321
16.7196 | -34.9509 | -26.4396
16.7257 | -34.9514 | -26.4469
16.7318 | -34.9519 | -26.4545
16.7379 | -34.9524 | -26.462
16.744 | -34.9529 | -26.4693
16.7501 | -34.9534 | -26.4767
16.7562 | -34.9539 | -26.4842
16.7623 | -34.9544 | -26.4918
16.7684 | -34.9549 | -26.4994
16.7745 | -34.9554 | -26.5069
16.7806 | -34.9559 | -26.5141
16.7867 | -34.9564 | -26.5217
16.7928 | -34.9569 | -26.5289
16.7989 | -34.9574 | -26.5363
16.805 | -34.9579 | -26.5438
16.8111 | -34.9584 | -26.5512
16.8172 | -34.9589 | -26.5586
16.8233 | -34.9594 | -26.5659
16.8294 | -34.9599 | -26.5734
16.8356 | -34.9604 | -26.5807
16.8417 | -34.9609 | -26.5882
16.8478 | -34.9613 | -26.5955
16.8539 | -34.9618 | -26.6031
16.86 | -34.9623 | -26.6105
16.8661 | -34.9628 | -26.618
16.8722 | -34.9633 | -26.6257
16.8783 | -34.9638 | -26.6332
16.8844 | -34.9643 | -26.6405
16.8905 | -34.9648 | -26.6479
16.8966 | -34.9653 | -26.6554
16.9027 | -34.9658 | -26.6628
16.9088 | -34.9663 | -26.6703
16.9149 | -34.9668 | -26.6777
16.921 | -34.9673 | -26.685
16.9271 | -34.9678 | -26.6924
16.9332 | -34.9683 | -26.6998
16.9393 | -34.9688 | -26.7071
16.9454 | -34.9692 | -26.7146
16.9515 | -34.9697 | -26.7221
16.9576 | -34.9703 | -26.7295
16.9637 | -34.9708 | -26.7368
16.9698 | -34.9713 | -26.7443
16.9759 | -34.9717 | -26.7518
16.9821 | -34.9722 | -26.7594
16.9882 | -34.9727 | -26.7667
16.9943 | -34.9732 | -26.7741
17.0004 | -34.9737 | -26.7816
17.0065 | -34.9742 | -26.789
17.0126 | -34.9746 | -26.7963
17.0187 | -34.9752 | -26.8038
17.0248 | -34.9757 | -26.8109
17.0309 | -34.9762 | -26.8184
17.037 | -34.9767 | -26.826
17.0431 | -34.9772 | -26.8334
17.0492 | -34.9777 | -26.8407
17.0553 | -34.9782 | -26.8481
17.0614 | -34.9787 | -26.8553
17.0675 | -34.9792 | -26.8627
17.0736 | -34.9796 | -26.8702
17.0797 | -34.9801 | -26.8777
17.0858 | -34.9807 | -26.8851
17.0919 | -34.9812 | -26.8925
17.098 | -34.9817 | -26.8994
17.1041 | -34.9822 | -26.9069
17.1102 | -34.9826 | -26.9142
17.1163 | -34.9831 | -26.9216
17.1225 | -34.9836 | -26.9292
17.1286 | -34.9841 | -26.9362
17.1347 | -34.9846 | -26.9434
17.1408 | -34.9851 | -26.9508
17.1469 | -34.9856 | -26.9581
17.153 | -34.9861 | -26.9655
17.1591 | -34.9866 | -26.9728
17.1652 | -34.9871 | -26.98
17.1713 | -34.9876 | -26.9874
17.1774 | -34.9881 | -26.9947
17.1835 | -34.9885 | -27.0023
17.1896 | -34.989 | -27.0097
17.1957 | -34.9896 | -27.017
17.2018 | -34.99 | -27.0242
17.2079 | -34.9905 | -27.0315
17.214 | -34.991 | -27.0388
17.2201 | -34.9915 | -27.0462
17.2262 | -34.9919 | -27.0536
17.2323 | -34.9924 | -27.061
17.2384 | -34.9929 | -27.068
17.2445 | -34.9934 | -27.0753
17.2506 | -34.9939 | -27.0826
17.2567 | -34.9945 | -27.09
17.2628 | -34.9949 | -27.0975
17.269 | -34.9954 | -27.1046
17.2751 | -34.9959 | -27.112
17.2812 | -34.9964 | -27.1193
17.2873 | -34.9969 | -27.1266
17.2934 | -34.9974 | -27.1339
17.2995 | -34.9979 | -27.1412
17.3056 | -34.9984 | -27.1484
17.3117 | -34.9989 | -27.1557
17.3178 | -34.9994 | -27.163
17.3239 | -34.9999 | -27.1703
17.33 | -35.0004 | -27.1777
17.3361 | -35.0009 | -27.1852
17.3422 | -35.0014 | -27.1924
17.3483 | -35.0019 | -27.1995
17.3544 | -35.0024 | -27.2069
17.3605 | -35.0029 | -27.2143
17.3666 | -35.0034 | -27.2216
17.3727 | -35.0039 | -27.2291
17.3788 | -35.0043 | -27.2362
17.3849 | -35.0048 | -27.2436
17.391 | -35.0053 | -27.2511
17.3971 | -35.0057 | -27.2583
17.4032 | -35.0062 | -27.2656
17.4094 | -35.0068 | -27.2725
17.4155 | -35.0073 | -27.2798
17.4216 | -35.0078 | -27.2872
17.4277 | -35.0083 | -27.2945
17.4338 | -35.0088 | -27.3019
17.4399 | -35.0093 | -27.3091
17.446 | -35.0098 | -27.3162
17.4521 | -35.0103 | -27.3235
17.4582 | -35.0108 | -27.3308
17.4643 | -35.0112 | -27.3381
17.4704 | -35.0117 | -27.3455
17.4765 | -35.0122 | -27.3527
17.4826 | -35.0128 | -27.3597
17.4887 | -35.0133 | -27.3671
17.4948 | -35.0137 | -27.3745
17.5009 | -35.0142 | -27.382
17.507 | -35.0147 | -27.3891
17.5131 | -35.0152 | -27.3962
17.5192 | -35.0157 | -27.4034
17.5253 | -35.0162 | -27.4106
17.5314 | -35.0167 | -27.4179
17.5375 | -35.0171 | -27.4252
17.5436 | -35.0176 | -27.4326
17.5497 | -35.0181 | -27.44
17.5559 | -35.0186 | -27.447
17.562 | -35.0191 | -27.4542
17.5681 | -35.0196 | -27.4615
17.5742 | -35.0201 | -27.4689
17.5803 | -35.0206 | -27.4761
17.5864 | -35.0211 | -27.4831
17.5925 | -35.0216 | -27.4904
17.5986 | -35.0221 | -27.4977
17.6047 | -35.0226 | -27.5048
17.6108 | -35.0231 | -27.512
17.6169 | -35.0236 | -27.5191
17.623 | -35.0241 | -27.5261
17.6291 | -35.0246 | -27.5333
17.6352 | -35.025 | -27.5406
17.6413 | -35.0255 | -27.5477
17.6474 | -35.026 | -27.555
17.6535 | -35.0265 | -27.5623
17.6596 | -35.027 | -27.5693
17.6657 | -35.0275 | -27.5765
17.6718 | -35.028 | -27.5838
17.6779 | -35.0285 | -27.5911
17.684 | -35.029 | -27.5984
17.6901 | -35.0295 | -27.6056
17.6963 | -35.03 | -27.6126
17.7024 | -35.0305 | -27.6197
17.7085 | -35.031 | -27.6269
17.7146 | -35.0315 | -27.6342
17.7207 | -35.0319 | -27.6413
17.7268 | -35.0324 | -27.6482
17.7329 | -35.0329 | -27.6554
17.739 | -35.0334 | -27.6626
17.7451 | -35.0339 | -27.6697
17.7512 | -35.0344 | -27.6769
17.7573 | -35.0349 | -27.684
17.7634 | -35.0354 | -27.691
17.7695 | -35.0359 | -27.698
17.7756 | -35.0364 | -27.7052
17.7817 | -35.0369 | -27.7123
17.7878 | -35.0374 | -27.7196
17.7939 | -35.0379 | -27.7268
17.8 | -35.0384 | -27.7339
17.8061 | -35.0389 | -27.7412
17.8122 | -35.0393 | -27.7483
17.8183 | -35.0398 | -27.7554
17.8244 | -35.0403 | -27.7624
17.8305 | -35.0408 | -27.7697
17.8366 | -35.0413 | -27.7765
17.8428 | -35.0418 | -27.7837
17.8489 | -35.0423 | -27.7908
17.855 | -35.0428 | -27.7979
17.8611 | -35.0432 | -27.8051
17.8672 | -35.0437 | -27.812
17.8733 | -35.0442 | -27.8192
17.8794 | -35.0447 | -27.8261
17.8855 | -35.0452 | -27.8334
17.8916 | -35.0457 | -27.8406
17.8977 | -35.0461 | -27.8477
17.9038 | -35.0466 | -27.8546
17.9099 | -35.0471 | -27.8618
17.916 | -35.0476 | -27.8689
17.9221 | -35.0481 | -27.8761
17.9282 | -35.0486 | -27.8832
17.9343 | -35.0491 | -27.8903
17.9404 | -35.0497 | -27.8973
17.9465 | -35.0501 | -27.9043
17.9526 | -35.0506 | -27.9114
17.9587 | -35.0511 | -27.9184
17.9648 | -35.0516 | -27.9255
17.9709 | -35.0521 | -27.9327
17.977 | -35.0526 | -27.9398
17.9832 | -35.0532 | -27.9469
17.9893 | -35.0537 | -27.9541
17.9954 | -35.0542 | -27.9613
18.0015 | -35.0546 | -27.9684
18.0076 | -35.0551 | -27.9753
18.0137 | -35.0556 | -27.9826
18.0198 | -35.0561 | -27.9896
18.0259 | -35.0566 | -27.9968
18.032 | -35.0571 | -28.0039
18.0381 | -35.0575 | -28.011
18.0442 | -35.0581 | -28.0179
18.0503 | -35.0586 | -28.0252
18.0564 | -35.0591 | -28.0324
18.0625 | -35.0596 | -28.0396
18.0686 | -35.0601 | -28.0468
18.0747 | -35.0606 | -28.0541
18.0808 | -35.0611 | -28.0609
18.0869 | -35.0615 | -28.0679
18.093 | -35.062 | -28.0751
18.0991 | -35.0625 | -28.082
18.1052 | -35.063 | -28.0893
18.1113 | -35.0635 | -28.0964
18.1174 | -35.064 | -28.1033
18.1236 | -35.0645 | -28.1104
18.1297 | -35.065 | -28.1175
18.1358 | -35.0655 | -28.1247
18.1419 | -35.066 | -28.1319
18.148 | -35.0665 | -28.1391
18.1541 | -35.067 | -28.1459
18.1602 | -35.0675 | -28.153
18.1663 | -35.068 | -28.1603
18.1724 | -35.0685 | -28.1672
18.1785 | -35.0689 | -28.1744
18.1846 | -35.0695 | -28.1813
18.1907 | -35.0699 | -28.1884
18.1968 | -35.0704 | -28.1953
18.2029 | -35.0709 | -28.2026
18.209 | -35.0714 | -28.2098
18.2151 | -35.0719 | -28.217
18.2212 | -35.0725 | -28.2238
18.2273 | -35.073 | -28.2308
18.2334 | -35.0735 | -28.2378
18.2395 | -35.074 | -28.2449
18.2456 | -35.0744 | -28.252
18.2517 | -35.075 | -28.2589
18.2578 | -35.0755 | -28.2656
18.2639 | -35.0759 | -28.2726
18.2701 | -35.0765 | -28.2796
18.2762 | -35.0769 | -28.2867
18.2823 | -35.0774 | -28.294
18.2884 | -35.0779 | -28.3011
18.2945 | -35.0784 | -28.3079
18.3006 | -35.0788 | -28.315
18.3067 | -35.0793 | -28.3222
18.3128 | -35.0798 | -28.3294
18.3189 | -35.0803 | -28.3362
18.325 | -35.0809 | -28.343
18.3311 | -35.0813 | -28.3501
18.3372 | -35.0818 | -28.3572
18.3433 | -35.0823 | -28.3643
18.3494 | -35.0828 | -28.3713
18.3555 | -35.0833 | -28.3783
18.3616 | -35.0838 | -28.3851
18.3677 | -35.0843 | -28.3922
18.3738 | -35.0848 | -28.3996
18.3799 | -35.0852 | -28.4067
18.386 | -35.0857 | -28.4137
18.3921 | -35.0862 | -28.4207
18.3982 | -35.0867 | -28.4274
18.4043 | -35.0872 | -28.4343
18.4105 | -35.0877 | -28.4414
18.4166 | -35.0882 | -28.4485
18.4227 | -35.0887 | -28.4556
18.4288 | -35.0892 | -28.4625
18.4349 | -35.0897 | -28.4693
18.441 | -35.0902 | -28.4764
18.4471 | -35.0906 | -28.4833
18.4532 | -35.0911 | -28.4901
18.4593 | -35.0916 | -28.4972
18.4654 | -35.0921 | -28.5042
18.4715 | -35.0926 | -28.5111
18.4776 | -35.0931 | -28.5181
18.4837 | -35.0936 | -28.5251
18.4898 | -35.0941 | -28.5324
18.4959 | -35.0946 | -28.5394
18.502 | -35.0951 | -28.5462
18.5081 | -35.0956 | -28.5533
18.5142 | -35.0961 | -28.5603
18.5203 | -35.0966 | -28.5673
18.5264 | -35.097 | -28.5743
18.5325 | -35.0975 | -28.5814
18.5386 | -35.0981 | -28.5881
18.5447 | -35.0985 | -28.5951
18.5508 | -35.099 | -28.6022
18.557 | -35.0995 | -28.6095
18.5631 | -35.1 | -28.6168
18.5692 | -35.1005 | -28.6241
18.5753 | -35.101 | -28.6308
18.5814 | -35.1015 | -28.6379
18.5875 | -35.102 | -28.6449
18.5936 | -35.1025 | -28.652
18.5997 | -35.103 | -28.659
18.6058 | -35.1035 | -28.6656
18.6119 | -35.104 | -28.6726
18.618 | -35.1045 | -28.6795
18.6241 | -35.105 | -28.6866
18.6302 | -35.1055 | -28.6936
18.6363 | -35.106 | -28.7008
18.6424 | -35.1065 | -28.7077
18.6485 | -35.1071 | -28.715
18.6546 | -35.1075 | -28.7221
18.6607 | -35.108 | -28.7291
18.6668 | -35.1085 | -28.7361
18.6729 | -35.109 | -28.7432
18.679 | -35.1095 | -28.75
18.6851 | -35.11 | -28.7568
18.6912 | -35.1105 | -28.7638
18.6974 | -35.111 | -28.7706
18.7035 | -35.1115 | -28.7774
18.7096 | -35.112 | -28.7844
18.7157 | -35.1125 | -28.7912
18.7218 | -35.113 | -28.798
18.7279 | -35.1135 | -28.8048
18.734 | -35.114 | -28.8119
18.7401 | -35.1145 | -28.8187
18.7462 | -35.115 | -28.8255
18.7523 | -35.1155 | -28.8325
18.7584 | -35.116 | -28.8393
18.7645 | -35.1165 | -28.8462
18.7706 | -35.117 | -28.8532
18.7767 | -35.1175 | -28.8601
18.7828 | -35.118 | -28.8669
18.7889 | -35.1185 | -28.8741
18.795 | -35.119 | -28.8812
18.8011 | -35.1195 | -28.8881
18.8072 | -35.12 | -28.8951
18.8133 | -35.1205 | -28.902
18.8194 | -35.1209 | -28.909
18.8255 | -35.1214 | -28.916
18.8316 | -35.1219 | -28.923
18.8377 | -35.1224 | -28.9301
18.8439 | -35.1229 | -28.9369
18.85 | -35.1234 | -28.9438
18.8561 | -35.1239 | -28.9505
18.8622 | -35.1244 | -28.9576
18.8683 | -35.1249 | -28.9645
18.8744 | -35.1253 | -28.9713
18.8805 | -35.1258 | -28.978
18.8866 | -35.1263 | -28.985
18.8927 | -35.1268 | -28.9919
18.8988 | -35.1273 | -28.9988
18.9049 | -35.1278 | -29.0055
18.911 | -35.1283 | -29.0126
18.9171 | -35.1288 | -29.0196
18.9232 | -35.1293 | -29.0262
18.9293 | -35.1299 | -29.0329
18.9354 | -35.1304 | -29.0398
18.9415 | -35.1309 | -29.0466
18.9476 | -35.1314 | -29.0538
18.9537 | -35.1319 | -29.0608
18.9598 | -35.1324 | -29.0676
18.9659 | -35.1328 | -29.0745
18.972 | -35.1334 | -29.0812
18.9781 | -35.1338 | -29.0882
18.9843 | -35.1344 | -29.0953
18.9904 | -35.1349 | -29.1021
18.9965 | -35.1354 | -29.1087
19.0026 | -35.1358 | -29.1153
19.0087 | -35.1363 | -29.1221
19.0148 | -35.1369 | -29.129
19.0209 | -35.1373 | -29.1361
19.027 | -35.1378 | -29.1427
19.0331 | -35.1383 | -29.1496
19.0392 | -35.1388 | -29.1565
19.0453 | -35.1392 | -29.1634
19.0514 | -35.1397 | -29.1701
19.0575 | -35.1402 | -29.1769
19.0636 | -35.1407 | -29.1837
19.0697 | -35.1412 | -29.1904
19.0758 | -35.1417 | -29.197
19.0819 | -35.1421 | -29.2038
19.088 | -35.1426 | -29.2105
19.0941 | -35.1431 | -29.2174
19.1002 | -35.1436 | -29.2243
19.1063 | -35.1441 | -29.2313
19.1124 | -35.1446 | -29.2383
19.1185 | -35.1451 | -29.2453
19.1246 | -35.1456 | -29.2521
19.1308 | -35.1461 | -29.2591
19.1369 | -35.1466 | -29.2657
19.143 | -35.1471 | -29.2727
19.1491 | -35.1476 | -29.2797
19.1552 | -35.1481 | -29.2865
19.1613 | -35.1486 | -29.2936
19.1674 | -35.1491 | -29.3007
19.1735 | -35.1496 | -29.3073
19.1796 | -35.1501 | -29.314
19.1857 | -35.1506 | -29.3208
19.1918 | -35.1511 | -29.3274
19.1979 | -35.1516 | -29.3344
19.204 | -35.1521 | -29.3408
19.2101 | -35.1526 | -29.3477
19.2162 | -35.1531 | -29.3545
19.2223 | -35.1536 | -29.3612
19.2284 | -35.154 | -29.3683
19.2345 | -35.1545 | -29.3751
19.2406 | -35.1551 | -29.3817
19.2467 | -35.1556 | -29.3887
19.2528 | -35.1561 | -29.3956
19.2589 | -35.1566 | -29.4025
19.265 | -35.1571 | -29.4093
19.2712 | -35.1575 | -29.4165
19.2773 | -35.158 | -29.4236
19.2834 | -35.1585 | -29.4306
19.2895 | -35.1589 | -29.4375
19.2956 | -35.1594 | -29.4445
19.3017 | -35.1599 | -29.4514
19.3078 | -35.1604 | -29.4583
19.3139 | -35.1609 | -29.465
19.32 | -35.1613 | -29.4718
19.3261 | -35.1618 | -29.4787
19.3322 | -35.1623 | -29.4854
19.3383 | -35.1628 | -29.4921
19.3444 | -35.1633 | -29.4988
19.3505 | -35.1638 | -29.5057
19.3566 | -35.1642 | -29.5122
19.3627 | -35.1647 | -29.519
19.3688 | -35.1653 | -29.5259
19.3749 | -35.1657 | -29.5325
19.381 | -35.1662 | -29.5392
19.3871 | -35.1667 | -29.5461
19.3932 | -35.1671 | -29.5527
19.3993 | -35.1676 | -29.5597
19.4054 | -35.1681 | -29.5666
19.4115 | -35.1685 | -29.5735
19.4177 | -35.169 | -29.5806
19.4238 | -35.1695 | -29.5874
19.4299 | -35.17 | -29.5943
19.436 | -35.1705 | -29.6011
19.4421 | -35.1709 | -29.6081
19.4482 | -35.1714 | -29.6152
19.4543 | -35.1719 | -29.6221
19.4604 | -35.1724 | -29.6292
19.4665 | -35.1729 | -29.6361
19.4726 | -35.1734 | -29.6425
19.4787 | -35.1739 | -29.6491
19.4848 | -35.1744 | -29.6561
19.4909 | -35.1748 | -29.663
19.497 | -35.1753 | -29.67
19.5031 | -35.1757 | -29.6768
19.5092 | -35.1763 | -29.6836
19.5153 | -35.1767 | -29.6903
19.5214 | -35.1772 | -29.6974
19.5275 | -35.1777 | -29.7044
19.5336 | -35.1782 | -29.7112
19.5397 | -35.1786 | -29.7179
19.5458 | -35.1791 | -29.7249
19.5519 | -35.1796 | -29.732
19.5581 | -35.18 | -29.7387
19.5642 | -35.1805 | -29.7457
19.5703 | -35.181 | -29.7527
19.5764 | -35.1815 | -29.7595
19.5825 | -35.182 | -29.7663
19.5886 | -35.1825 | -29.7732
19.5947 | -35.183 | -29.78
19.6008 | -35.1834 | -29.7869
19.6069 | -35.1839 | -29.7935
19.613 | -35.1844 | -29.8002
19.6191 | -35.185 | -29.8072
19.6252 | -35.1854 | -29.8141
19.6313 | -35.1859 | -29.8207
19.6374 | -35.1863 | -29.8277
19.6435 | -35.1868 | -29.8344
19.6496 | -35.1872 | -29.8413
19.6557 | -35.1877 | -29.8482
19.6618 | -35.1882 | -29.8546
19.6679 | -35.1887 | -29.8616
19.674 | -35.1892 | -29.8684
19.6801 | -35.1897 | -29.8753
19.6862 | -35.1902 | -29.8823
19.6923 | -35.1907 | -29.8889
19.6984 | -35.1911 | -29.8957
19.7046 | -35.1916 | -29.9025
19.7107 | -35.1921 | -29.9095
19.7168 | -35.1925 | -29.9165
19.7229 | -35.193 | -29.9233
19.729 | -35.1935 | -29.9303
19.7351 | -35.194 | -29.9371
19.7412 | -35.1945 | -29.944
19.7473 | -35.195 | -29.9508
19.7534 | -35.1955 | -29.9577
19.7595 | -35.196 | -29.9646
19.7656 | -35.1964 | -29.9714
19.7717 | -35.1969 | -29.978
19.7778 | -35.1974 | -29.9848
19.7839 | -35.1979 | -29.9916
19.79 | -35.1983 | -29.9985
19.7961 | -35.1988 | -30.0054
19.8022 | -35.1993 | -30.0124
19.8083 | -35.1998 | -30.019
19.8144 | -35.2003 | -30.0258
19.8205 | -35.2007 | -30.0327
19.8266 | -35.2012 | -30.0394
19.8327 | -35.2017 | -30.0463
19.8388 | -35.2022 | -30.0528
19.845 | -35.2026 | -30.0598
19.8511 | -35.2031 | -30.0666
19.8572 | -35.2036 | -30.0736
19.8633 | -35.2041 | -30.0804
19.8694 | -35.2046 | -30.0872
19.8755 | -35.2051 | -30.0937
19.8816 | -35.2056 | -30.1006
19.8877 | -35.2061 | -30.1076
19.8938 | -35.2065 | -30.1142
19.8999 | -35.207 | -30.1207
19.906 | -35.2075 | -30.1276
19.9121 | -35.208 | -30.1341
19.9182 | -35.2085 | -30.1409
19.9243 | -35.209 | -30.1478
19.9304 | -35.2095 | -30.1546
19.9365 | -35.21 | -30.1615
19.9426 | -35.2104 | -30.1684
19.9487 | -35.2109 | -30.175
19.9548 | -35.2114 | -30.182
19.9609 | -35.2119 | -30.1886
19.967 | -35.2123 | -30.1955
19.9731 | -35.2128 | -30.2023
19.9792 | -35.2133 | -30.2092
19.9853 | -35.2138 | -30.2158
19.9915 | -35.2142 | -30.2225
19.9976 | -35.2147 | -30.2294
20.0037 | -35.2152 | -30.2362
20.0098 | -35.2157 | -30.2429
20.0159 | -35.2162 | -30.2495
20.022 | -35.2166 | -30.2562
20.0281 | -35.2171 | -30.2632
20.0342 | -35.2176 | -30.2699
20.0403 | -35.218 | -30.2769
20.0464 | -35.2185 | -30.2836
20.0525 | -35.219 | -30.2904
20.0586 | -35.2195 | -30.2971
20.0647 | -35.22 | -30.3038
20.0708 | -35.2204 | -30.3108
20.0769 | -35.2209 | -30.3177
20.083 | -35.2214 | -30.3244
20.0891 | -35.2219 | -30.3312
20.0952 | -35.2224 | -30.3378
20.1013 | -35.2228 | -30.3447
20.1074 | -35.2233 | -30.3516
20.1135 | -35.2238 | -30.3584
20.1196 | -35.2243 | -30.3652
20.1257 | -35.2248 | -30.372
20.1319 | -35.2252 | -30.379
20.138 | -35.2257 | -30.3856
20.1441 | -35.2262 | -30.3925
20.1502 | -35.2266 | -30.3996
20.1563 | -35.2271 | -30.4064
20.1624 | -35.2275 | -30.4132
20.1685 | -35.228 | -30.42
20.1746 | -35.2285 | -30.4267
20.1807 | -35.2289 | -30.4333
20.1868 | -35.2294 | -30.4402
20.1929 | -35.2299 | -30.4469
20.199 | -35.2304 | -30.4538
20.2051 | -35.2309 | -30.4605
20.2112 | -35.2313 | -30.4674
20.2173 | -35.2318 | -30.4742
20.2234 | -35.2323 | -30.4809
20.2295 | -35.2328 | -30.4877
20.2356 | -35.2333 | -30.4944
20.2417 | -35.2337 | -30.5011
20.2478 | -35.2342 | -30.5079
20.2539 | -35.2347 | -30.5146
20.26 | -35.2352 | -30.5213
20.2661 | -35.2357 | -30.5281
20.2723 | -35.2361 | -30.5351
20.2784 | -35.2366 | -30.542
20.2845 | -35.2371 | -30.5489
20.2906 | -35.2375 | -30.5556
20.2967 | -35.238 | -30.5622
20.3028 | -35.2385 | -30.5691
20.3089 | -35.2389 | -30.576
20.315 | -35.2394 | -30.5828
20.3211 | -35.2399 | -30.5895
20.3272 | -35.2404 | -30.5962
20.3333 | -35.2408 | -30.6027
20.3394 | -35.2413 | -30.6095
20.3455 | -35.2418 | -30.6162
20.3516 | -35.2422 | -30.623
20.3577 | -35.2427 | -30.6298
20.3638 | -35.2431 | -30.6365
20.3699 | -35.2436 | -30.6431
20.376 | -35.2441 | -30.6499
20.3821 | -35.2446 | -30.6568
20.3882 | -35.2451 | -30.6636
20.3943 | -35.2456 | -30.6705
20.4004 | -35.246 | -30.6772
20.4065 | -35.2464 | -30.6841
20.4126 | -35.2469 | -30.691
20.4188 | -35.2474 | -30.6979
20.4249 | -35.2479 | -30.7045
20.431 | -35.2483 | -30.7112
20.4371 | -35.2487 | -30.7178
20.4432 | -35.2492 | -30.7246
20.4493 | -35.2496 | -30.7313
20.4554 | -35.2501 | -30.738
20.4615 | -35.2506 | -30.7447
20.4676 | -35.2511 | -30.7515
20.4737 | -35.2516 | -30.7582
20.4798 | -35.2521 | -30.7649
20.4859 | -35.2526 | -30.7716
20.492 | -35.253 | -30.7782
20.4981 | -35.2534 | -30.785
20.5042 | -35.2539 | -30.7919
20.5103 | -35.2544 | -30.7984
20.5164 | -35.2549 | -30.8051
20.5225 | -35.2553 | -30.812
20.5286 | -35.2558 | -30.8189
20.5347 | -35.2563 | -30.8256
20.5408 | -35.2567 | -30.8322
20.5469 | -35.2572 | -30.8389
20.553 | -35.2577 | -30.8457
20.5592 | -35.2581 | -30.8524
20.5653 | -35.2586 | -30.8592
20.5714 | -35.2591 | -30.866
20.5775 | -35.2595 | -30.8724
20.5836 | -35.26 | -30.8792
20.5897 | -35.2605 | -30.8859
20.5958 | -35.261 | -30.8927
20.6019 | -35.2614 | -30.8994
20.608 | -35.2619 | -30.9064
20.6141 | -35.2624 | -30.913
20.6202 | -35.2629 | -30.9195
20.6263 | -35.2633 | -30.926
20.6324 | -35.2638 | -30.9328
20.6385 | -35.2642 | -30.9396
20.6446 | -35.2647 | -30.9463
20.6507 | -35.2651 | -30.9527
20.6568 | -35.2656 | -30.9594
20.6629 | -35.2661 | -30.966
20.669 | -35.2665 | -30.9728
20.6751 | -35.267 | -30.9798
20.6812 | -35.2674 | -30.9865
20.6873 | -35.2679 | -30.9933
20.6934 | -35.2684 | -31.0004
20.6995 | -35.2688 | -31.0075
20.7057 | -35.2693 | -31.0141
20.7118 | -35.2698 | -31.0209
20.7179 | -35.2703 | -31.0274
20.724 | -35.2707 | -31.034
20.7301 | -35.2712 | -31.0409
20.7362 | -35.2716 | -31.0475
20.7423 | -35.2721 | -31.0543
20.7484 | -35.2725 | -31.061
20.7545 | -35.273 | -31.0671
20.7606 | -35.2735 | -31.0738
20.7667 | -35.2739 | -31.0802
20.7728 | -35.2744 | -31.0869
20.7789 | -35.2748 | -31.0934
20.785 | -35.2753 | -31.1
20.7911 | -35.2758 | -31.1065
20.7972 | -35.2762 | -31.1131
20.8033 | -35.2767 | -31.1201
20.8094 | -35.2772 | -31.1266
20.8155 | -35.2777 | -31.1334
20.8216 | -35.2781 | -31.1402
20.8277 | -35.2785 | -31.1471
20.8338 | -35.279 | -31.1537
20.8399 | -35.2795 | -31.1605
20.8461 | -35.28 | -31.167
20.8522 | -35.2804 | -31.1736
20.8583 | -35.2808 | -31.1798
20.8644 | -35.2813 | -31.1867
20.8705 | -35.2818 | -31.1931
20.8766 | -35.2822 | -31.2002
20.8827 | -35.2827 | -31.2072
20.8888 | -35.2831 | -31.2137
20.8949 | -35.2836 | -31.2203
20.901 | -35.284 | -31.227
20.9071 | -35.2845 | -31.2335
20.9132 | -35.2849 | -31.2401
20.9193 | -35.2854 | -31.247
20.9254 | -35.2859 | -31.2537
20.9315 | -35.2863 | -31.2602
20.9376 | -35.2867 | -31.2669
20.9437 | -35.2872 | -31.2734
20.9498 | -35.2877 | -31.2803
20.9559 | -35.2881 | -31.2872
20.962 | -35.2886 | -31.2938
20.9681 | -35.289 | -31.3005
20.9742 | -35.2895 | -31.307
20.9803 | -35.29 | -31.3137
20.9864 | -35.2904 | -31.3203
20.9926 | -35.2909 | -31.327
20.9987 | -35.2913 | -31.3333
21.0048 | -35.2918 | -31.3401
21.0109 | -35.2922 | -31.3466
21.017 | -35.2927 | -31.3533
21.0231 | -35.2932 | -31.36
21.0292 | -35.2936 | -31.3665
21.0353 | -35.294 | -31.3729
21.0414 | -35.2944 | -31.3796
21.0475 | -35.2949 | -31.3863
21.0536 | -35.2954 | -31.3929
21.0597 | -35.2958 | -31.3997
21.0658 | -35.2962 | -31.4063
21.0719 | -35.2967 | -31.4123
21.078 | -35.2972 | -31.4189
21.0841 | -35.2976 | -31.4258
21.0902 | -35.2981 | -31.4325
21.0963 | -35.2985 | -31.4391
21.1024 | -35.299 | -31.4457
21.1085 | -35.2995 | -31.452
21.1146 | -35.2999 | -31.4587
21.1207 | -35.3004 | -31.4654
21.1268 | -35.3008 | -31.4723
21.133 | -35.3013 | -31.4789
21.1391 | -35.3018 | -31.4857
21.1452 | -35.3023 | -31.4923
21.1513 | -35.3027 | -31.499
21.1574 | -35.3032 | -31.5056
21.1635 | -35.3036 | -31.5123
21.1696 | -35.3041 | -31.5188
21.1757 | -35.3045 | -31.5254
21.1818 | -35.3049 | -31.5318
21.1879 | -35.3054 | -31.5387
21.194 | -35.3059 | -31.5455
21.2001 | -35.3063 | -31.5518
21.2062 | -35.3068 | -31.5586
21.2123 | -35.3073 | -31.5654
21.2184 | -35.3078 | -31.5718
21.2245 | -35.3082 | -31.5789
21.2306 | -35.3086 | -31.5859
21.2367 | -35.3091 | -31.5924
21.2428 | -35.3095 | -31.5984
21.2489 | -35.3101 | -31.6045
21.255 | -35.3106 | -31.6111
21.2611 | -35.311 | -31.6179
21.2672 | -35.3115 | -31.6248
21.2733 | -35.3119 | -31.6316
21.2795 | -35.3124 | -31.6376
21.2856 | -35.3128 | -31.6443
21.2917 | -35.3133 | -31.6509
21.2978 | -35.3137 | -31.6576
21.3039 | -35.3142 | -31.664
21.31 | -35.3146 | -31.6706
21.3161 | -35.315 | -31.6771
21.3222 | -35.3155 | -31.6834
21.3283 | -35.3159 | -31.6902
21.3344 | -35.3163 | -31.6971
21.3405 | -35.3168 | -31.7037
21.3466 | -35.3173 | -31.7104
21.3527 | -35.3177 | -31.7168
21.3588 | -35.3182 | -31.7232
21.3649 | -35.3187 | -31.7296
21.371 | -35.3192 | -31.7364
21.3771 | -35.3196 | -31.7431
21.3832 | -35.32 | -31.7497
21.3893 | -35.3205 | -31.7561
21.3954 | -35.3209 | -31.7629
21.4015 | -35.3214 | -31.7695
21.4076 | -35.3218 | -31.7763
21.4137 | -35.3223 | -31.7829
21.4199 | -35.3227 | -31.7895
21.426 | -35.3232 | -31.796
21.4321 | -35.3236 | -31.8026
21.4382 | -35.324 | -31.809
21.4443 | -35.3245 | -31.8155
21.4504 | -35.325 | -31.8219
21.4565 | -35.3254 | -31.8283
21.4626 | -35.3258 | -31.8346
21.4687 | -35.3263 | -31.8412
21.4748 | -35.3267 | -31.8477
21.4809 | -35.3272 | -31.8543
21.487 | -35.3276 | -31.8607
21.4931 | -35.3281 | -31.8672
21.4992 | -35.3285 | -31.8738
21.5053 | -35.3289 | -31.8804
21.5114 | -35.3294 | -31.8871
21.5175 | -35.3299 | -31.8937
21.5236 | -35.3303 | -31.9002
21.5297 | -35.3307 | -31.9066
21.5358 | -35.3312 | -31.913
21.5419 | -35.3316 | -31.9194
21.548 | -35.3321 | -31.9259
21.5541 | -35.3325 | -31.9324
21.5602 | -35.333 | -31.9388
21.5664 | -35.3334 | -31.9451
21.5725 | -35.3339 | -31.9514
21.5786 | -35.3343 | -31.9579
21.5847 | -35.3348 | -31.9649
21.5908 | -35.3352 | -31.9716
21.5969 | -35.3357 | -31.9779
21.603 | -35.3361 | -31.9845
21.6091 | -35.3366 | -31.9911
21.6152 | -35.3371 | -31.9976
21.6213 | -35.3375 | -32.0041
21.6274 | -35.3379 | -32.0105
21.6335 | -35.3384 | -32.017
21.6396 | -35.3389 | -32.0236
21.6457 | -35.3394 | -32.0301
21.6518 | -35.3398 | -32.0363
21.6579 | -35.3403 | -32.0429
21.664 | -35.3408 | -32.0498
21.6701 | -35.3412 | -32.0563
21.6762 | -35.3417 | -32.0627
21.6823 | -35.3421 | -32.0691
21.6884 | -35.3426 | -32.0753
21.6945 | -35.343 | -32.0819
21.7006 | -35.3435 | -32.0886
21.7068 | -35.3439 | -32.0951
21.7129 | -35.3444 | -32.1014
21.719 | -35.3448 | -32.1079
21.7251 | -35.3452 | -32.1144
21.7312 | -35.3457 | -32.1208
21.7373 | -35.3462 | -32.127
21.7434 | -35.3466 | -32.1335
21.7495 | -35.347 | -32.14
21.7556 | -35.3475 | -32.1466
21.7617 | -35.3479 | -32.1535
21.7678 | -35.3484 | -32.1603
21.7739 | -35.3488 | -32.1665
21.78 | -35.3493 | -32.173
21.7861 | -35.3498 | -32.1796
21.7922 | -35.3502 | -32.1862
21.7983 | -35.3506 | -32.1928
21.8044 | -35.3511 | -32.1993
21.8105 | -35.3516 | -32.2056
21.8166 | -35.352 | -32.2121
21.8227 | -35.3525 | -32.2184
21.8288 | -35.353 | -32.225
21.8349 | -35.3534 | -32.2318
21.841 | -35.3539 | -32.2382
21.8471 | -35.3544 | -32.2442
21.8533 | -35.3548 | -32.2507
21.8594 | -35.3552 | -32.2571
21.8655 | -35.3557 | -32.2635
21.8716 | -35.3561 | -32.27
21.8777 | -35.3566 | -32.2763
21.8838 | -35.357 | -32.2826
21.8899 | -35.3574 | -32.2892
21.896 | -35.3579 | -32.2956
21.9021 | -35.3583 | -32.302
21.9082 | -35.3588 | -32.3085
21.9143 | -35.3592 | -32.3145
21.9204 | -35.3597 | -32.3207
21.9265 | -35.3602 | -32.327
21.9326 | -35.3606 | -32.3333
21.9387 | -35.3611 | -32.3398
21.9448 | -35.3616 | -32.3464
21.9509 | -35.362 | -32.3525
21.957 | -35.3625 | -32.359
21.9631 | -35.3629 | -32.3653
21.9692 | -35.3633 | -32.372
21.9753 | -35.3638 | -32.3784
21.9814 | -35.3642 | -32.385
21.9875 | -35.3647 | -32.3907
21.9937 | -35.3652 | -32.3973
21.9998 | -35.3657 | -32.4035
22.0059 | -35.3661 | -32.4102
22.012 | -35.3665 | -32.4165
22.0181 | -35.367 | -32.4223
22.0242 | -35.3674 | -32.4286
22.0303 | -35.3678 | -32.4352
22.0364 | -35.3683 | -32.4419
22.0425 | -35.3687 | -32.4484
22.0486 | -35.3692 | -32.4548
22.0547 | -35.3696 | -32.4605
22.0608 | -35.37 | -32.4666
22.0669 | -35.3705 | -32.473
22.073 | -35.3709 | -32.4796
22.0791 | -35.3714 | -32.4859
22.0852 | -35.3718 | -32.4925
22.0913 | -35.3723 | -32.4985
22.0974 | -35.3727 | -32.505
22.1035 | -35.3732 | -32.5115
22.1096 | -35.3737 | -32.518
22.1157 | -35.3742 | -32.5246
22.1218 | -35.3746 | -32.5311
22.1279 | -35.3751 | -32.5374
22.134 | -35.3756 | -32.5437
22.1402 | -35.376 | -32.5499
22.1463 | -35.3764 | -32.5561
22.1524 | -35.3769 | -32.5624
22.1585 | -35.3774 | -32.5688
22.1646 | -35.3778 | -32.5752
22.1707 | -35.3782 | -32.5815
22.1768 | -35.3787 | -32.5879
22.1829 | -35.3791 | -32.5942
22.189 | -35.3796 | -32.6004
22.1951 | -35.3801 | -32.6066
22.2012 | -35.3805 | -32.613
22.2073 | -35.381 | -32.6195
22.2134 | -35.3814 | -32.6259
22.2195 | -35.3819 | -32.6323
22.2256 | -35.3824 | -32.6386
22.2317 | -35.3827 | -32.6445
22.2378 | -35.3832 | -32.6507
22.2439 | -35.3837 | -32.6572
22.25 | -35.3841 | -32.6639
22.2561 | -35.3845 | -32.67
22.2622 | -35.385 | -32.6764
22.2683 | -35.3855 | -32.6822
22.2744 | -35.3859 | -32.6887
22.2806 | -35.3864 | -32.6949
22.2867 | -35.3868 | -32.7012
22.2928 | -35.3873 | -32.7075
22.2989 | -35.3877 | -32.7137
22.305 | -35.3882 | -32.7196
22.3111 | -35.3887 | -32.7259
22.3172 | -35.3891 | -32.7324
22.3233 | -35.3895 | -32.7386
22.3294 | -35.39 | -32.7449
22.3355 | -35.3905 | -32.7508
22.3416 | -35.3909 | -32.7572
22.3477 | -35.3914 | -32.7635
22.3538 | -35.3918 | -32.7699
22.3599 | -35.3922 | -32.776
22.366 | -35.3927 | -32.7822
22.3721 | -35.3932 | -32.7882
22.3782 | -35.3936 | -32.7945
22.3843 | -35.3941 | -32.8007
22.3904 | -35.3945 | -32.8069
22.3965 | -35.395 | -32.8131
22.4026 | -35.3954 | -32.8193
22.4087 | -35.3959 | -32.826
22.4148 | -35.3963 | -32.8326
22.4209 | -35.3968 | -32.8387
22.4271 | -35.3972 | -32.8452
22.4332 | -35.3977 | -32.8515
22.4393 | -35.3981 | -32.8577
22.4454 | -35.3986 | -32.8635
22.4515 | -35.399 | -32.8701
22.4576 | -35.3995 | -32.8762
22.4637 | -35.3999 | -32.8823
22.4698 | -35.4004 | -32.8884
22.4759 | -35.4008 | -32.8942
22.482 | -35.4013 | -32.9
22.4881 | -35.4017 | -32.9061
22.4942 | -35.4021 | -32.9123
22.5003 | -35.4025 | -32.9185
22.5064 | -35.403 | -32.9247
22.5125 | -35.4035 | -32.9299
22.5186 | -35.404 | -32.9361
22.5247 | -35.4044 | -32.9424
22.5308 | -35.4049 | -32.9489
22.5369 | -35.4054 | -32.9554
22.543 | -35.4058 | -32.9619
22.5491 | -35.4062 | -32.9679
22.5552 | -35.4066 | -32.974
22.5613 | -35.4071 | -32.98
22.5675 | -35.4075 | -32.9863
22.5736 | -35.408 | -32.9923
22.5797 | -35.4084 | -32.9983
22.5858 | -35.4089 | -33.0039
22.5919 | -35.4094 | -33.0102
22.598 | -35.4098 | -33.0165
22.6041 | -35.4102 | -33.0225
22.6102 | -35.4107 | -33.0285
22.6163 | -35.4112 | -33.0343
22.6224 | -35.4116 | -33.0407
22.6285 | -35.4121 | -33.0469
22.6346 | -35.4125 | -33.0534
22.6407 | -35.413 | -33.0593
22.6468 | -35.4134 | -33.0653
22.6529 | -35.4139 | -33.071
22.659 | -35.4144 | -33.0771
22.6651 | -35.4149 | -33.0832
22.6712 | -35.4153 | -33.0896
22.6773 | -35.4157 | -33.0957
22.6834 | -35.4162 | -33.1019
22.6895 | -35.4166 | -33.108
22.6956 | -35.4171 | -33.1143
22.7017 | -35.4176 | -33.1205
22.7079 | -35.418 | -33.1267
22.714 | -35.4185 | -33.133
22.7201 | -35.4189 | -33.1392
22.7262 | -35.4194 | -33.1453
22.7323 | -35.4198 | -33.1513
22.7384 | -35.4202 | -33.1574
22.7445 | -35.4207 | -33.1632
22.7506 | -35.4211 | -33.1695
22.7567 | -35.4215 | -33.1757
22.7628 | -35.4219 | -33.1817
22.7689 | -35.4224 | -33.1876
22.775 | -35.4228 | -33.1939
22.7811 | -35.4232 | -33.1999
22.7872 | -35.4236 | -33.2061
22.7933 | -35.4241 | -33.2118
22.7994 | -35.4246 | -33.2179
22.8055 | -35.425 | -33.2239
22.8116 | -35.4254 | -33.2298
22.8177 | -35.4259 | -33.2359
22.8238 | -35.4264 | -33.2419
22.8299 | -35.4268 | -33.2473
22.836 | -35.4273 | -33.2534
22.8421 | -35.4277 | -33.2593
22.8482 | -35.4282 | -33.2653
22.8544 | -35.4286 | -33.2716
22.8605 | -35.429 | -33.2778
22.8666 | -35.4295 | -33.2838
22.8727 | -35.4299 | -33.2899
22.8788 | -35.4304 | -33.2959
22.8849 | -35.4308 | -33.3021
22.891 | -35.4312 | -33.3083
22.8971 | -35.4317 | -33.3142
22.9032 | -35.4321 | -33.3198
22.9093 | -35.4326 | -33.3258
22.9154 | -35.433 | -33.3315
22.9215 | -35.4335 | -33.3376
22.9276 | -35.4339 | -33.3437
22.9337 | -35.4344 | -33.3496
22.9398 | -35.4349 | -33.3556
22.9459 | -35.4354 | -33.3616
22.952 | -35.4358 | -33.3678
22.9581 | -35.4362 | -33.3737
22.9642 | -35.4367 | -33.38
22.9703 | -35.4372 | -33.3859
22.9764 | -35.4376 | -33.3922
22.9825 | -35.4381 | -33.398
22.9886 | -35.4385 | -33.4037
22.9948 | -35.4389 | -33.41
23.0009 | -35.4394 | -33.4163
23.007 | -35.4399 | -33.4216
23.0131 | -35.4403 | -33.4276
23.0192 | -35.4408 | -33.4338
23.0253 | -35.4413 | -33.4396
23.0314 | -35.4417 | -33.4457
23.0375 | -35.4422 | -33.4516
23.0436 | -35.4426 | -33.4575
23.0497 | -35.4431 | -33.4636
23.0558 | -35.4435 | -33.4696
23.0619 | -35.444 | -33.4759
23.068 | -35.4444 | -33.4819
23.0741 | -35.4449 | -33.4876
23.0802 | -35.4453 | -33.4934
23.0863 | -35.4458 | -33.4996
23.0924 | -35.4463 | -33.5054
23.0985 | -35.4467 | -33.5113
23.1046 | -35.4471 | -33.5173
23.1107 | -35.4476 | -33.523
23.1168 | -35.448 | -33.5289
23.1229 | -35.4485 | -33.535
23.129 | -35.4489 | -33.5409
23.1351 | -35.4494 | -33.547
23.1413 | -35.4499 | -33.5528
23.1474 | -35.4503 | -33.5584
23.1535 | -35.4508 | -33.5643
23.1596 | -35.4513 | -33.5703
23.1657 | -35.4518 | -33.5763
23.1718 | -35.4522 | -33.5825
23.1779 | -35.4527 | -33.5886
23.184 | -35.4531 | -33.5942
23.1901 | -35.4536 | -33.6001
23.1962 | -35.454 | -33.6061
23.2023 | -35.4545 | -33.612
23.2084 | -35.455 | -33.618
23.2145 | -35.4554 | -33.6241
23.2206 | -35.4558 | -33.6304
23.2267 | -35.4562 | -33.6362
23.2328 | -35.4567 | -33.6422
23.2389 | -35.4571 | -33.6484
23.245 | -35.4576 | -33.6542
23.2511 | -35.458 | -33.6601
23.2572 | -35.4585 | -33.6662
23.2633 | -35.4589 | -33.6721
23.2694 | -35.4594 | -33.6782
23.2755 | -35.4598 | -33.6842
23.2817 | -35.4602 | -33.6903
23.2878 | -35.4607 | -33.6958
23.2939 | -35.4612 | -33.7016
23.3 | -35.4616 | -33.7076
23.3061 | -35.462 | -33.7134
23.3122 | -35.4625 | -33.7189
23.3183 | -35.463 | -33.7248
23.3244 | -35.4635 | -33.7304
23.3305 | -35.4639 | -33.7365
23.3366 | -35.4645 | -33.7421
23.3427 | -35.4649 | -33.7478
23.3488 | -35.4654 | -33.7538
23.3549 | -35.4658 | -33.7598
23.361 | -35.4663 | -33.765
23.3671 | -35.4667 | -33.7711
23.3732 | -35.4672 | -33.7767
23.3793 | -35.4676 | -33.7827
23.3854 | -35.468 | -33.7889
23.3915 | -35.4685 | -33.7944
23.3976 | -35.4689 | -33.8001
23.4037 | -35.4694 | -33.806
23.4098 | -35.4698 | -33.812
23.4159 | -35.4702 | -33.8182
23.422 | -35.4706 | -33.8242
23.4282 | -35.4711 | -33.83
23.4343 | -35.4716 | -33.8359
23.4404 | -35.472 | -33.8415
23.4465 | -35.4724 | -33.8472
23.4526 | -35.4729 | -33.8533
23.4587 | -35.4733 | -33.8591
23.4648 | -35.4738 | -33.8648
23.4709 | -35.4742 | -33.871
23.477 | -35.4747 | -33.877
23.4831 | -35.4751 | -33.8829
23.4892 | -35.4756 | -33.8889
23.4953 | -35.476 | -33.8943
23.5014 | -35.4766 | -33.8999
23.5075 | -35.477 | -33.9054
23.5136 | -35.4775 | -33.9114
23.5197 | -35.478 | -33.9173
23.5258 | -35.4784 | -33.9233
23.5319 | -35.4789 | -33.9289
23.538 | -35.4793 | -33.9346
23.5441 | -35.4798 | -33.9405
23.5502 | -35.4802 | -33.9464
23.5563 | -35.4807 | -33.9519
23.5624 | -35.4811 | -33.958
23.5686 | -35.4816 | -33.9634
23.5747 | -35.4821 | -33.9693
23.5808 | -35.4825 | -33.9754
23.5869 | -35.4829 | -33.9811
23.593 | -35.4834 | -33.9869
23.5991 | -35.4838 | -33.9927
23.6052 | -35.4843 | -33.9983
23.6113 | -35.4848 | -34.0044
23.6174 | -35.4852 | -34.0103
23.6235 | -35.4856 | -34.0163
23.6296 | -35.4861 | -34.0222
23.6357 | -35.4865 | -34.0278
23.6418 | -35.487 | -34.0334
23.6479 | -35.4874 | -34.0392
23.654 | -35.4878 | -34.0449
23.6601 | -35.4883 | -34.0508
23.6662 | -35.4887 | -34.0567
23.6723 | -35.4892 | -34.0624
23.6784 | -35.4896 | -34.0682
23.6845 | -35.49 | -34.074
23.6906 | -35.4905 | -34.0796
23.6967 | -35.4909 | -34.0854
23.7028 | -35.4913 | -34.091
23.7089 | -35.4919 | -34.0971
23.7151 | -35.4923 | -34.1029
23.7212 | -35.4928 | -34.1084
23.7273 | -35.4932 | -34.1144
23.7334 | -35.4937 | -34.1201
23.7395 | -35.4941 | -34.1259
23.7456 | -35.4945 | -34.1312
23.7517 | -35.495 | -34.1369
23.7578 | -35.4954 | -34.1425
23.7639 | -35.4959 | -34.1482
23.77 | -35.4963 | -34.1541
23.7761 | -35.4968 | -34.1598
23.7822 | -35.4972 | -34.1655
23.7883 | -35.4976 | -34.1711
23.7944 | -35.4981 | -34.1767
23.8005 | -35.4985 | -34.1823
23.8066 | -35.4989 | -34.1882
23.8127 | -35.4993 | -34.1936
23.8188 | -35.4998 | -34.1993
23.8249 | -35.5002 | -34.205
23.831 | -35.5007 | -34.2109
23.8371 | -35.5011 | -34.2168
23.8432 | -35.5015 | -34.2224
23.8493 | -35.502 | -34.2276
23.8555 | -35.5024 | -34.2332
23.8616 | -35.5029 | -34.2392
23.8677 | -35.5033 | -34.2455
23.8738 | -35.5038 | -34.2511
23.8799 | -35.5042 | -34.2567
23.886 | -35.5047 | -34.2621
23.8921 | -35.5051 | -34.2678
23.8982 | -35.5056 | -34.2736
23.9043 | -35.506 | -34.2794
23.9104 | -35.5065 | -34.2851
23.9165 | -35.5069 | -34.2909
23.9226 | -35.5074 | -34.2962
23.9287 | -35.5078 | -34.3024
23.9348 | -35.5083 | -34.3081
23.9409 | -35.5088 | -34.3139
23.947 | -35.5092 | -34.3196
23.9531 | -35.5097 | -34.325
23.9592 | -35.5101 | -34.3309
23.9653 | -35.5106 | -34.3365
23.9714 | -35.511 | -34.3422
23.9775 | -35.5115 | -34.348
23.9836 | -35.512 | -34.3536
23.9897 | -35.5125 | -34.3591
23.9958 | -35.5129 | -34.3647
24.002 | -35.5134 | -34.3705
24.0081 | -35.5138 | -34.3764
24.0142 | -35.5142 | -34.3821
24.0203 | -35.5147 | -34.3878
24.0264 | -35.5152 | -34.3934
24.0325 | -35.5156 | -34.3992
24.0386 | -35.5161 | -34.4049
24.0447 | -35.5165 | -34.4107
24.0508 | -35.5169 | -34.4165
24.0569 | -35.5174 | -34.4221
24.063 | -35.5178 | -34.4275
24.0691 | -35.5183 | -34.4334
24.0752 | -35.5188 | -34.4389
24.0813 | -35.5192 | -34.4449
24.0874 | -35.5197 | -34.4505
24.0935 | -35.5201 | -34.4564
24.0996 | -35.5206 | -34.4621
24.1057 | -35.5211 | -34.4682
24.1118 | -35.5215 | -34.474
24.1179 | -35.522 | -34.4802
24.124 | -35.5224 | -34.4856
24.1301 | -35.5229 | -34.4908
24.1362 | -35.5233 | -34.4968
24.1424 | -35.5238 | -34.5023
24.1485 | -35.5243 | -34.5077
24.1546 | -35.5247 | -34.5132
24.1607 | -35.5252 | -34.519
24.1668 | -35.5256 | -34.5245
24.1729 | -35.5261 | -34.5303
24.179 | -35.5266 | -34.5357
24.1851 | -35.527 | -34.5414
24.1912 | -35.5275 | -34.5472
24.1973 | -35.5279 | -34.5529
24.2034 | -35.5283 | -34.5585
24.2095 | -35.5288 | -34.564
24.2156 | -35.5292 | -34.5697
24.2217 | -35.5297 | -34.5754
24.2278 | -35.5301 | -34.5811
24.2339 | -35.5305 | -34.5867
24.24 | -35.5309 | -34.5923
24.2461 | -35.5314 | -34.5978
24.2522 | -35.5318 | -34.6035
24.2583 | -35.5323 | -34.609
24.2644 | -35.5328 | -34.6146
24.2705 | -35.5332 | -34.62
24.2766 | -35.5336 | -34.6254
24.2827 | -35.5341 | -34.6311
24.2889 | -35.5345 | -34.6371
24.295 | -35.5349 | -34.6426
24.3011 | -35.5353 | -34.6487
24.3072 | -35.5358 | -34.6539
24.3133 | -35.5363 | -34.6598
24.3194 | -35.5367 | -34.6656
24.3255 | -35.5371 | -34.6715
24.3316 | -35.5376 | -34.6773
24.3377 | -35.538 | -34.6831
24.3438 | -35.5385 | -34.6884
24.3499 | -35.5389 | -34.6941
24.356 | -35.5394 | -34.6998
24.3621 | -35.5398 | -34.7056
24.3682 | -35.5402 | -34.711
24.3743 | -35.5406 | -34.7164
24.3804 | -35.5411 | -34.7216
24.3865 | -35.5415 | -34.727
24.3926 | -35.542 | -34.7326
24.3987 | -35.5424 | -34.7383
24.4048 | -35.5428 | -34.7437
24.4109 | -35.5431 | -34.749
24.417 | -35.5436 | -34.7546
24.4231 | -35.544 | -34.7603
24.4293 | -35.5444 | -34.7657
24.4354 | -35.5449 | -34.771
24.4415 | -35.5453 | -34.7765
24.4476 | -35.5457 | -34.7823
24.4537 | -35.5461 | -34.7877
24.4598 | -35.5465 | -34.7935
24.4659 | -35.5469 | -34.7987
24.472 | -35.5473 | -34.8044
24.4781 | -35.5477 | -34.8101
24.4842 | -35.5481 | -34.8155
24.4903 | -35.5486 | -34.8209
24.4964 | -35.549 | -34.8268
24.5025 | -35.5494 | -34.8326
24.5086 | -35.5498 | -34.8383
24.5147 | -35.5502 | -34.8438
24.5208 | -35.5506 | -34.8493
24.5269 | -35.551 | -34.8547
24.533 | -35.5515 | -34.8604
24.5391 | -35.5519 | -34.8658
24.5452 | -35.5524 | -34.8713
24.5513 | -35.5528 | -34.877
24.5574 | -35.5533 | -34.8825
24.5635 | -35.5537 | -34.8882
24.5696 | -35.5542 | -34.8938
24.5758 | -35.5546 | -34.8993
24.5819 | -35.555 | -34.9048
24.588 | -35.5555 | -34.9105
24.5941 | -35.556 | -34.9163
24.6002 | -35.5564 | -34.9216
24.6063 | -35.5569 | -34.9271
24.6124 | -35.5573 | -34.9327
24.6185 | -35.5577 | -34.9381
24.6246 | -35.5582 | -34.9432
24.6307 | -35.5586 | -34.9488
24.6368 | -35.5591 | -34.9544
24.6429 | -35.5595 | -34.9601
24.649 | -35.56 | -34.9656
24.6551 | -35.5604 | -34.9715
24.6612 | -35.5609 | -34.9765
24.6673 | -35.5613 | -34.9825
24.6734 | -35.5618 | -34.9881
24.6795 | -35.5622 | -34.9936
24.6856 | -35.5627 | -34.9994
24.6917 | -35.5632 | -35.0049
24.6978 | -35.5636 | -35.0108
24.7039 | -35.564 | -35.0164
24.71 | -35.5644 | -35.0219
24.7162 | -35.5649 | -35.0272
24.7223 | -35.5653 | -35.0329
24.7284 | -35.5657 | -35.0386
24.7345 | -35.5662 | -35.0442
24.7406 | -35.5666 | -35.05
24.7467 | -35.5671 | -35.0556
24.7528 | -35.5675 | -35.0608
24.7589 | -35.568 | -35.0662
24.765 | -35.5684 | -35.0715
24.7711 | -35.5688 | -35.0772
24.7772 | -35.5693 | -35.0826
24.7833 | -35.5697 | -35.0883
24.7894 | -35.5701 | -35.0941
24.7955 | -35.5706 | -35.0996
24.8016 | -35.5709 | -35.1047
24.8077 | -35.5714 | -35.1103
24.8138 | -35.5718 | -35.1156
24.8199 | -35.5722 | -35.1211
24.826 | -35.5727 | -35.1268
24.8321 | -35.5732 | -35.1328
24.8382 | -35.5736 | -35.1384
24.8443 | -35.5741 | -35.1442
24.8504 | -35.5745 | -35.1497
24.8565 | -35.5749 | -35.1555
24.8627 | -35.5753 | -35.1611
24.8688 | -35.5757 | -35.1667
24.8749 | -35.5761 | -35.1723
24.881 | -35.5766 | -35.1778
24.8871 | -35.577 | -35.1835
24.8932 | -35.5775 | -35.1888
24.8993 | -35.5779 | -35.1945
24.9054 | -35.5783 | -35.1993
24.9115 | -35.5787 | -35.2051
24.9176 | -35.5792 | -35.2104
24.9237 | -35.5797 | -35.2159
24.9298 | -35.5802 | -35.2215
24.9359 | -35.5806 | -35.2274
24.942 | -35.581 | -35.2323
24.9481 | -35.5814 | -35.2378
24.9542 | -35.5819 | -35.2431
24.9603 | -35.5823 | -35.2484
24.9664 | -35.5827 | -35.254
24.9725 | -35.5831 | -35.2593
24.9786 | -35.5835 | -35.2644
24.9847 | -35.5839 | -35.2697
24.9908 | -35.5843 | -35.2753
24.9969 | -35.5847 | -35.2806
25.0031 | -35.5851 | -35.2863
25.0092 | -35.5856 | -35.2918
25.0153 | -35.586 | -35.2973
25.0214 | -35.5864 | -35.3032
25.0275 | -35.5869 | -35.309
25.0336 | -35.5873 | -35.3143
25.0397 | -35.5877 | -35.3199
25.0458 | -35.5881 | -35.3251
25.0519 | -35.5886 | -35.3307
25.058 | -35.589 | -35.3363
25.0641 | -35.5894 | -35.3419
25.0702 | -35.5898 | -35.3472
25.0763 | -35.5902 | -35.3525
25.0824 | -35.5906 | -35.3579
25.0885 | -35.591 | -35.3633
25.0946 | -35.5914 | -35.3689
25.1007 | -35.5918 | -35.3742
25.1068 | -35.5923 | -35.3796
25.1129 | -35.5927 | -35.3852
25.119 | -35.5931 | -35.3906
25.1251 | -35.5935 | -35.396
25.1312 | -35.5939 | -35.4013
25.1373 | -35.5943 | -35.4066
25.1435 | -35.5947 | -35.4121
25.1496 | -35.5951 | -35.4176
25.1557 | -35.5955 | -35.423
25.1618 | -35.5959 | -35.4284
25.1679 | -35.5964 | -35.434
25.174 | -35.5968 | -35.4395
25.1801 | -35.5972 | -35.4452
25.1862 | -35.5976 | -35.4507
25.1923 | -35.598 | -35.4565
25.1984 | -35.5985 | -35.4618
25.2045 | -35.5989 | -35.4672
25.2106 | -35.5994 | -35.4728
25.2167 | -35.5997 | -35.4782
25.2228 | -35.6001 | -35.4836
25.2289 | -35.6005 | -35.4892
25.235 | -35.601 | -35.4946
25.2411 | -35.6014 | -35.5
25.2472 | -35.6018 | -35.5055
25.2533 | -35.6022 | -35.5113
25.2594 | -35.6026 | -35.5165
25.2655 | -35.603 | -35.5222
25.2716 | -35.6034 | -35.5276
25.2777 | -35.6038 | -35.5327
25.2838 | -35.6042 | -35.5383
25.29 | -35.6046 | -35.5438
25.2961 | -35.605 | -35.5488
25.3022 | -35.6054 | -35.5542
25.3083 | -35.6058 | -35.5597
25.3144 | -35.6062 | -35.5651
25.3205 | -35.6066 | -35.5707
25.3266 | -35.607 | -35.5761
25.3327 | -35.6075 | -35.5815
25.3388 | -35.6079 | -35.5871
25.3449 | -35.6083 | -35.5926
25.351 | -35.6087 | -35.5981
25.3571 | -35.6091 | -35.6036
25.3632 | -35.6095 | -35.6091
25.3693 | -35.6099 | -35.6145
25.3754 | -35.6103 | -35.6196
25.3815 | -35.6108 | -35.625
25.3876 | -35.6112 | -35.6305
25.3937 | -35.6116 | -35.636
25.3998 | -35.6119 | -35.6412
25.4059 | -35.6123 | -35.647
25.412 | -35.6127 | -35.6521
25.4181 | -35.6131 | -35.6574
25.4242 | -35.6135 | -35.663
25.4304 | -35.6139 | -35.6685
25.4365 | -35.6143 | -35.6734
25.4426 | -35.6148 | -35.6787
25.4487 | -35.6151 | -35.6844
25.4548 | -35.6155 | -35.6898
25.4609 | -35.6159 | -35.6952
25.467 | -35.6163 | -35.7004
25.4731 | -35.6168 | -35.7058
25.4792 | -35.6172 | -35.7112
25.4853 | -35.6176 | -35.7165
25.4914 | -35.618 | -35.722
25.4975 | -35.6184 | -35.7273
25.5036 | -35.6187 | -35.7327
25.5097 | -35.6191 | -35.7382
25.5158 | -35.6196 | -35.7437
25.5219 | -35.62 | -35.7492
25.528 | -35.6204 | -35.7545
25.5341 | -35.6208 | -35.7598
25.5402 | -35.6212 | -35.7654
25.5463 | -35.6216 | -35.7706
25.5524 | -35.622 | -35.7762
25.5585 | -35.6224 | -35.7813
25.5646 | -35.6228 | -35.7865
25.5707 | -35.6232 | -35.792
25.5769 | -35.6236 | -35.7974
25.583 | -35.624 | -35.8028
25.5891 | -35.6244 | -35.8083
25.5952 | -35.6249 | -35.814
25.6013 | -35.6253 | -35.8194
25.6074 | -35.6257 | -35.8246
25.6135 | -35.6261 | -35.8298
25.6196 | -35.6265 | -35.8352
25.6257 | -35.6269 | -35.8404
25.6318 | -35.6273 | -35.8459
25.6379 | -35.6277 | -35.851
25.644 | -35.6281 | -35.8564
25.6501 | -35.6285 | -35.8615
25.6562 | -35.6289 | -35.867
25.6623 | -35.6293 | -35.8723
25.6684 | -35.6297 | -35.8777
25.6745 | -35.6301 | -35.883
25.6806 | -35.6306 | -35.8883
25.6867 | -35.631 | -35.8939
25.6928 | -35.6314 | -35.8997
25.6989 | -35.6318 | -35.905
25.705 | -35.6322 | -35.9103
25.7111 | -35.6326 | -35.9156
25.7173 | -35.633 | -35.9208
25.7234 | -35.6335 | -35.9259
25.7295 | -35.6339 | -35.9314
25.7356 | -35.6343 | -35.9369
25.7417 | -35.6347 | -35.9422
25.7478 | -35.635 | -35.9475
25.7539 | -35.6354 | -35.9528
25.76 | -35.6358 | -35.958
25.7661 | -35.6361 | -35.9634
25.7722 | -35.6366 | -35.9688
25.7783 | -35.637 | -35.9738
25.7844 | -35.6374 | -35.979
25.7905 | -35.6377 | -35.9843
25.7966 | -35.6381 | -35.9899
25.8027 | -35.6385 | -35.9951
25.8088 | -35.6389 | -36.0004
25.8149 | -35.6393 | -36.0059
25.821 | -35.6396 | -36.0109
25.8271 | -35.64 | -36.0161
25.8332 | -35.6405 | -36.0216
25.8393 | -35.6408 | -36.027
25.8454 | -35.6412 | -36.0322
25.8515 | -35.6416 | -36.038
25.8576 | -35.642 | -36.0431
25.8638 | -35.6424 | -36.0484
25.8699 | -35.6428 | -36.0537
25.876 | -35.6433 | -36.059
25.8821 | -35.6436 | -36.0646
25.8882 | -35.644 | -36.0699
25.8943 | -35.6444 | -36.0752
25.9004 | -35.6448 | -36.0808
25.9065 | -35.6452 | -36.0863
25.9126 | -35.6456 | -36.0917
25.9187 | -35.6461 | -36.0969
25.9248 | -35.6464 | -36.1021
25.9309 | -35.6468 | -36.1078
25.937 | -35.6472 | -36.1133
25.9431 | -35.6476 | -36.1188
25.9492 | -35.648 | -36.1245
25.9553 | -35.6484 | -36.1299
25.9614 | -35.6488 | -36.1357
25.9675 | -35.6492 | -36.141
25.9736 | -35.6496 | -36.1463
25.9797 | -35.65 | -36.1518
25.9858 | -35.6504 | -36.1571
25.9919 | -35.6508 | -36.1624
25.998 | -35.6512 | -36.1677
26.0042 | -35.6516 | -36.1734
26.0103 | -35.652 | -36.1788
26.0164 | -35.6524 | -36.1838
26.0225 | -35.6528 | -36.1891
26.0286 | -35.6532 | -36.1944
26.0347 | -35.6535 | -36.1994
26.0408 | -35.6538 | -36.2048
26.0469 | -35.6543 | -36.2101
26.053 | -35.6546 | -36.2155
26.0591 | -35.655 | -36.2209
26.0652 | -35.6554 | -36.2262
26.0713 | -35.6558 | -36.2318
26.0774 | -35.6561 | -36.2376
26.0835 | -35.6565 | -36.2431
26.0896 | -35.6569 | -36.2482
26.0957 | -35.6573 | -36.2533
26.1018 | -35.6576 | -36.2583
26.1079 | -35.658 | -36.2637
26.114 | -35.6584 | -36.269
26.1201 | -35.6588 | -36.2744
26.1262 | -35.6592 | -36.2797
26.1323 | -35.6596 | -36.2849
26.1384 | -35.6599 | -36.2899
26.1445 | -35.6603 | -36.2955
26.1507 | -35.6606 | -36.3011
26.1568 | -35.661 | -36.3064
26.1629 | -35.6615 | -36.3118
26.169 | -35.6619 | -36.3173
26.1751 | -35.6623 | -36.3223
26.1812 | -35.6626 | -36.3278
26.1873 | -35.663 | -36.3334
26.1934 | -35.6634 | -36.3385
26.1995 | -35.6638 | -36.3439
26.2056 | -35.6642 | -36.3489
26.2117 | -35.6646 | -36.3545
26.2178 | -35.6649 | -36.36
26.2239 | -35.6653 | -36.3649
26.23 | -35.6656 | -36.3702
26.2361 | -35.666 | -36.3754
26.2422 | -35.6664 | -36.3806
26.2483 | -35.6668 | -36.3864
26.2544 | -35.6671 | -36.3918
26.2605 | -35.6675 | -36.3973
26.2666 | -35.6679 | -36.4026
26.2727 | -35.6683 | -36.4078
26.2788 | -35.6686 | -36.4133
26.2849 | -35.669 | -36.4189
26.2911 | -35.6694 | -36.4246
26.2972 | -35.6698 | -36.4296
26.3033 | -35.6702 | -36.4349
26.3094 | -35.6706 | -36.4404
26.3155 | -35.6709 | -36.4457
26.3216 | -35.6714 | -36.4516
26.3277 | -35.6718 | -36.457
26.3338 | -35.6722 | -36.4625
26.3399 | -35.6725 | -36.468
26.346 | -35.673 | -36.4729
26.3521 | -35.6734 | -36.4785
26.3582 | -35.6738 | -36.4839
26.3643 | -35.6741 | -36.4892
26.3704 | -35.6745 | -36.4943
26.3765 | -35.6749 | -36.4995
26.3826 | -35.6752 | -36.505
26.3887 | -35.6756 | -36.5104
26.3948 | -35.676 | -36.5159
26.4009 | -35.6764 | -36.5213
26.407 | -35.6768 | -36.5266
26.4131 | -35.6771 | -36.532
26.4192 | -35.6775 | -36.537
26.4253 | -35.6779 | -36.5423
26.4314 | -35.6782 | -36.5474
26.4376 | -35.6786 | -36.5531
26.4437 | -35.679 | -36.5588
26.4498 | -35.6794 | -36.5641
26.4559 | -35.6797 | -36.569
26.462 | -35.6801 | -36.5744
26.4681 | -35.6805 | -36.5796
26.4742 | -35.6808 | -36.5848
26.4803 | -35.6812 | -36.5898
26.4864 | -35.6816 | -36.5949
26.4925 | -35.682 | -36.5999
26.4986 | -35.6824 | -36.6049
26.5047 | -35.6827 | -36.6101
26.5108 | -35.6831 | -36.6155
26.5169 | -35.6835 | -36.6206
26.523 | -35.6839 | -36.6255
26.5291 | -35.6843 | -36.6306
26.5352 | -35.6847 | -36.6359
26.5413 | -35.6851 | -36.6413
26.5474 | -35.6854 | -36.6469
26.5535 | -35.6858 | -36.6521
26.5596 | -35.6862 | -36.6569
26.5657 | -35.6866 | -36.6622
26.5718 | -35.687 | -36.6678
26.578 | -35.6874 | -36.673
26.5841 | -35.6878 | -36.6782
26.5902 | -35.6881 | -36.6839
26.5963 | -35.6885 | -36.6888
26.6024 | -35.6889 | -36.6939
26.6085 | -35.6892 | -36.6994
26.6146 | -35.6896 | -36.7043
26.6207 | -35.69 | -36.7097
26.6268 | -35.6903 | -36.7147
26.6329 | -35.6907 | -36.7197
26.639 | -35.6911 | -36.725
26.6451 | -35.6914 | -36.7303
26.6512 | -35.6918 | -36.7357
26.6573 | -35.6922 | -36.7411
26.6634 | -35.6925 | -36.7464
26.6695 | -35.6929 | -36.7519
26.6756 | -35.6933 | -36.7571
26.6817 | -35.6936 | -36.7627
26.6878 | -35.694 | -36.7675
26.6939 | -35.6944 | -36.773
26.7 | -35.6948 | -36.7784
26.7061 | -35.6952 | -36.7836
26.7122 | -35.6956 | -36.7886
26.7183 | -35.6959 | -36.7939
26.7245 | -35.6964 | -36.7989
26.7306 | -35.6967 | -36.8037
26.7367 | -35.6971 | -36.8084
26.7428 | -35.6975 | -36.8142
26.7489 | -35.6978 | -36.8192
26.755 | -35.6982 | -36.8242
26.7611 | -35.6986 | -36.8294
26.7672 | -35.6989 | -36.8348
26.7733 | -35.6993 | -36.8399
26.7794 | -35.6996 | -36.8452
26.7855 | -35.7001 | -36.8506
26.7916 | -35.7004 | -36.8562
26.7977 | -35.7008 | -36.8613
26.8038 | -35.7011 | -36.8666
26.8099 | -35.7015 | -36.8722
26.816 | -35.7018 | -36.878
26.8221 | -35.7022 | -36.8836
26.8282 | -35.7026 | -36.8892
26.8343 | -35.7029 | -36.8944
26.8404 | -35.7033 | -36.8995
26.8465 | -35.7036 | -36.9049
26.8526 | -35.704 | -36.91
26.8587 | -35.7044 | -36.9156
26.8649 | -35.7047 | -36.9207
26.871 | -35.7051 | -36.9258
26.8771 | -35.7055 | -36.9306
26.8832 | -35.7058 | -36.9358
26.8893 | -35.7062 | -36.9414
26.8954 | -35.7065 | -36.9467
26.9015 | -35.7068 | -36.9522
26.9076 | -35.7072 | -36.957
26.9137 | -35.7075 | -36.9626
26.9198 | -35.7079 | -36.9677
26.9259 | -35.7082 | -36.9727
26.932 | -35.7086 | -36.9778
26.9381 | -35.709 | -36.9829
26.9442 | -35.7093 | -36.988
26.9503 | -35.7096 | -36.9928
26.9564 | -35.71 | -36.998
26.9625 | -35.7103 | -37.0029
26.9686 | -35.7106 | -37.0082
26.9747 | -35.711 | -37.0136
26.9808 | -35.7112 | -37.0188
26.9869 | -35.7116 | -37.0241
26.993 | -35.712 | -37.0294
26.9991 | -35.7123 | -37.0348
27.0052 | -35.7127 | -37.04
27.0114 | -35.7131 | -37.0457
27.0175 | -35.7134 | -37.0505
27.0236 | -35.7137 | -37.0554
27.0297 | -35.7141 | -37.0606
27.0358 | -35.7145 | -37.0658
27.0419 | -35.7148 | -37.0709
27.048 | -35.7152 | -37.0764
27.0541 | -35.7156 | -37.0814
27.0602 | -35.7159 | -37.0866
27.0663 | -35.7163 | -37.0921
27.0724 | -35.7166 | -37.0973
27.0785 | -35.717 | -37.1027
27.0846 | -35.7174 | -37.108
27.0907 | -35.7178 | -37.113
27.0968 | -35.7182 | -37.1183
27.1029 | -35.7185 | -37.1234
27.109 | -35.7188 | -37.1286
27.1151 | -35.7192 | -37.1339
27.1212 | -35.7196 | -37.1391
27.1273 | -35.7199 | -37.1441
27.1334 | -35.7202 | -37.1491
27.1395 | -35.7207 | -37.1544
27.1456 | -35.721 | -37.1597
27.1518 | -35.7213 | -37.1653
27.1579 | -35.7216 | -37.1703
27.164 | -35.722 | -37.1756
27.1701 | -35.7224 | -37.1808
27.1762 | -35.7227 | -37.1859
27.1823 | -35.723 | -37.1914
27.1884 | -35.7234 | -37.197
27.1945 | -35.7238 | -37.2022
27.2006 | -35.7241 | -37.2074
27.2067 | -35.7245 | -37.2122
27.2128 | -35.7249 | -37.2174
27.2189 | -35.7252 | -37.223
27.225 | -35.7256 | -37.2279
27.2311 | -35.7259 | -37.233
27.2372 | -35.7263 | -37.2384
27.2433 | -35.7267 | -37.2434
27.2494 | -35.727 | -37.2485
27.2555 | -35.7273 | -37.2535
27.2616 | -35.7277 | -37.2588
27.2677 | -35.7281 | -37.2642
27.2738 | -35.7284 | -37.2694
27.2799 | -35.7288 | -37.2745
27.286 | -35.7291 | -37.2794
27.2921 | -35.7295 | -37.2844
27.2983 | -35.7299 | -37.2894
27.3044 | -35.7303 | -37.2947
27.3105 | -35.7306 | -37.2999
27.3166 | -35.731 | -37.3049
27.3227 | -35.7313 | -37.3104
27.3288 | -35.7317 | -37.3157
27.3349 | -35.732 | -37.3206
27.341 | -35.7324 | -37.3263
27.3471 | -35.7327 | -37.3314
27.3532 | -35.733 | -37.3364
27.3593 | -35.7334 | -37.3414
27.3654 | -35.7338 | -37.3465
27.3715 | -35.7341 | -37.3512
27.3776 | -35.7345 | -37.356
27.3837 | -35.7348 | -37.361
27.3898 | -35.7351 | -37.3661
27.3959 | -35.7356 | -37.3714
27.402 | -35.7359 | -37.376
27.4081 | -35.7362 | -37.3809
27.4142 | -35.7366 | -37.3862
27.4203 | -35.7369 | -37.3916
27.4264 | -35.7372 | -37.3968
27.4325 | -35.7376 | -37.4018
27.4387 | -35.738 | -37.4073
27.4448 | -35.7384 | -37.4128
27.4509 | -35.7388 | -37.4182
27.457 | -35.7392 | -37.4231
27.4631 | -35.7395 | -37.4281
27.4692 | -35.74 | -37.4333
27.4753 | -35.7403 | -37.4384
27.4814 | -35.7406 | -37.4436
27.4875 | -35.741 | -37.4484
27.4936 | -35.7413 | -37.4537
27.4997 | -35.7417 | -37.4588
27.5058 | -35.742 | -37.4634
27.5119 | -35.7424 | -37.4685
27.518 | -35.7428 | -37.4734
27.5241 | -35.7431 | -37.479
27.5302 | -35.7434 | -37.4842
27.5363 | -35.7438 | -37.4893
27.5424 | -35.7441 | -37.4941
27.5485 | -35.7444 | -37.4991
27.5546 | -35.7447 | -37.5043
27.5607 | -35.745 | -37.5097
27.5668 | -35.7454 | -37.5144
27.5729 | -35.7458 | -37.5194
27.5791 | -35.7461 | -37.5246
27.5852 | -35.7465 | -37.5295
27.5913 | -35.7468 | -37.5343
27.5974 | -35.7471 | -37.5398
27.6035 | -35.7475 | -37.5444
27.6096 | -35.7478 | -37.5499
27.6157 | -35.7481 | -37.5546
27.6218 | -35.7485 | -37.5593
27.6279 | -35.7489 | -37.5645
27.634 | -35.7492 | -37.5699
27.6401 | -35.7497 | -37.575
27.6462 | -35.75 | -37.5805
27.6523 | -35.7503 | -37.5854
27.6584 | -35.7507 | -37.5906
27.6645 | -35.751 | -37.5958
27.6706 | -35.7513 | -37.601
27.6767 | -35.7516 | -37.6062
27.6828 | -35.7519 | -37.6108
27.6889 | -35.7523 | -37.6162
27.695 | -35.7526 | -37.6212
27.7011 | -35.753 | -37.6261
27.7072 | -35.7533 | -37.6311
27.7133 | -35.7537 | -37.6364
27.7194 | -35.7541 | -37.6408
27.7256 | -35.7544 | -37.6464
27.7317 | -35.7548 | -37.6516
27.7378 | -35.7551 | -37.6567
27.7439 | -35.7555 | -37.6618
27.75 | -35.7559 | -37.6672
27.7561 | -35.7561 | -37.6719
27.7622 | -35.7565 | -37.6769
27.7683 | -35.7568 | -37.6821
27.7744 | -35.7572 | -37.6875
27.7805 | -35.7576 | -37.6923
27.7866 | -35.758 | -37.6974
27.7927 | -35.7583 | -37.7025
27.7988 | -35.7586 | -37.7077
27.8049 | -35.759 | -37.7128
27.811 | -35.7593 | -37.7172
27.8171 | -35.7597 | -37.7222
27.8232 | -35.7599 | -37.7267
27.8293 | -35.7603 | -37.7322
27.8354 | -35.7606 | -37.7376
27.8415 | -35.761 | -37.7422
27.8476 | -35.7613 | -37.747
27.8537 | -35.7616 | -37.7515
27.8598 | -35.762 | -37.757
27.866 | -35.7624 | -37.7616
27.8721 | -35.7627 | -37.767
27.8782 | -35.7631 | -37.7719
27.8843 | -35.7635 | -37.7773
27.8904 | -35.7637 | -37.783
27.8965 | -35.764 | -37.7875
27.9026 | -35.7644 | -37.7924
27.9087 | -35.7648 | -37.7975
27.9148 | -35.7651 | -37.8026
27.9209 | -35.7655 | -37.8074
27.927 | -35.7659 | -37.812
27.9331 | -35.7663 | -37.8171
27.9392 | -35.7666 | -37.8221
27.9453 | -35.7669 | -37.8269
27.9514 | -35.7671 | -37.8324
27.9575 | -35.7675 | -37.8378
27.9636 | -35.7679 | -37.8429
27.9697 | -35.7681 | -37.8472
27.9758 | -35.7684 | -37.8521
27.9819 | -35.7687 | -37.8573
27.988 | -35.7691 | -37.8623
27.9941 | -35.7694 | -37.8677
28.0002 | -35.7696 | -37.8729
28.0063 | -35.7698 | -37.8784
28.0125 | -35.7701 | -37.8837
28.0186 | -35.7705 | -37.889
28.0247 | -35.7709 | -37.8937
28.0308 | -35.7712 | -37.8985
28.0369 | -35.7715 | -37.9039
28.043 | -35.7718 | -37.9093
28.0491 | -35.7722 | -37.9138
28.0552 | -35.7725 | -37.9186
28.0613 | -35.7728 | -37.924
28.0674 | -35.773 | -37.9292
28.0735 | -35.7735 | -37.9333

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5576 | -42.8209
21.4908 | -35.581 | -35.0874
16.4948 | -34.9479 | -27.6957
12.6602 | -34.5787 | -21.4209
9.7171 | -34.3597 | -16.5117
7.4582 | -34.206 | -12.5264
5.7244 | -34.0996 | -9.3006
4.3936 | -34.0235 | -6.7014
3.3722 | -34.0124 | -4.651
2.5883 | -34.0174 | -3.0928
1.9866 | -34.0097 | -1.852
1.5248 | -34.0174 | -0.8559
1.1703 | -34.0277 | -0.20645
1 | -34.031 | 0.1163
0.89822 | -34.0219 | 0.38634
0.68941 | -34.013 | 0.96746
0.52919 | -34.0107 | 1.4172
0.40612 | -34.0148 | 2.0228
0.31171 | -34.0091 | 2.4716
0.23928 | -33.9988 | 3.1974
0.18358 | -34.0318 | 3.6085
0.14091 | -34.0275 | 4.4895
0.10819 | -34.069 | 5.6389
0.083017 | -34.0789 | 7.1088
0.063705 | -34.1074 | 8.9718
0.048923 | -34.1883 | 11.4439
0.037479 | -34.2939 | 14.6106
0.028801 | -34.4655 | 18.6041
0.022078 | -34.771 | 23.9277
0.017023 | -35.2479 | 30.3078
0.013018 | -36.0434 | 37.957
0.010014 | -37.1519 | 45.8364

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003673 | -0.82402
0.007347 | -24.233
0.01102 | -33.1679
0.014693 | -40.9896
0.018366 | -44.5003
0.02204 | -47.3382
0.025713 | -49.3212
0.029386 | -49.5648
0.033059 | -50.1136
0.036733 | -51.7123
0.040406 | -53.3175
0.044079 | -53.9421
0.047752 | -54.9799
0.051426 | -56.0245
0.055099 | -57.2554
0.058772 | -57.2713
0.062445 | -57.8318
0.066119 | -58.7566
0.069792 | -59.6318
0.073465 | -60.5176
0.077139 | -61.3215
0.080812 | -62.4337
0.084485 | -61.7331
0.088158 | -62.0937
0.091832 | -62.7046
0.095505 | -63.2928
0.099178 | -63.316
0.10285 | -64.0287
0.10652 | -64.7249
0.1102 | -64.2726
0.11387 | -64.3902
0.11754 | -64.9515
0.12122 | -66.0002
0.12489 | -65.64
0.12856 | -65.8629
0.13224 | -66.0392
0.13591 | -67.094
0.13958 | -68.7032
0.14326 | -68.1458
0.14693 | -67.716
0.1506 | -67.3958
0.15428 | -68.0262
0.15795 | -69.5006
0.16162 | -69.9881
0.1653 | -70.7048
0.16897 | -70.2709
0.17264 | -70.0399
0.17632 | -70.829
0.17999 | -70.7394
0.18366 | -70.2126
0.18734 | -70.1245
0.19101 | -69.3108
0.19468 | -69.0825
0.19836 | -70.3365
0.20203 | -70.3984
0.2057 | -71.1934
0.20938 | -71.9902
0.21305 | -72.5784
0.21672 | -72.5265
0.2204 | -72.2001
0.22407 | -72.1566
0.22774 | -72.0261
0.23142 | -71.8203
0.23509 | -72.5596
0.23876 | -72.7443
0.24244 | -72.2546
0.24611 | -72.7118
0.24978 | -74.202
0.25345 | -74.6729
0.25713 | -74.3834
0.2608 | -74.0638
0.26448 | -73.8742
0.26815 | -74.1705
0.27182 | -74.554
0.27549 | -74.449
0.27917 | -74.9692
0.28284 | -74.6428
0.28652 | -75.6039
0.29019 | -76.2734
0.29386 | -75.3149
0.29753 | -75.4246
0.30121 | -75.5462
0.30488 | -75.5322
0.30855 | -75.4288
0.31223 | -76.2382
0.3159 | -76.3724
0.31957 | -76.3244
0.32325 | -76.33
0.32692 | -76.0956
0.33059 | -76.9239
0.33427 | -76.5501
0.33794 | -76.8787
0.34161 | -75.9582
0.34529 | -75.7767
0.34896 | -76.9426
0.35263 | -77.2988
0.35631 | -76.7494
0.35998 | -76.5563
0.36365 | -76.9065
0.36733 | -76.7969
0.371 | -76.675
0.37467 | -76.5343
0.37835 | -77.2643
0.38202 | -77.9126
0.38569 | -79.222
0.38937 | -79.7437
0.39304 | -78.5993
0.39671 | -77.766
0.40039 | -77.9374
0.40406 | -78.3406
0.40773 | -78.3516
0.41141 | -79.0253
0.41508 | -78.891
0.41875 | -78.0495
0.42242 | -78.7401
0.4261 | -79.3213
0.42977 | -79.298
0.43345 | -80.6145
0.43712 | -80.6415
0.44079 | -80.7987
0.44446 | -80.2502
0.44814 | -79.6839
0.45181 | -79.3358
0.45548 | -80.0892
0.45916 | -79.8105
0.46283 | -79.3491
0.4665 | -79.6461
0.47018 | -79.861
0.47385 | -80.0898
0.47752 | -79.388
0.4812 | -80.0101
0.48487 | -80.5435
0.48854 | -80.4761
0.49222 | -81.2463
0.49589 | -80.948
0.49956 | -80.5238
0.50324 | -81.2114
0.50691 | -81.7511
0.51058 | -80.4438
0.51426 | -79.3382
0.51793 | -79.9164
0.5216 | -81.2541
0.52528 | -82.8755
0.52895 | -82.6596
0.53262 | -82.7569
0.5363 | -82.4515
0.53997 | -81.8063
0.54364 | -81.4513
0.54732 | -81.3902
0.55099 | -81.5437
0.55466 | -81.729
0.55834 | -81.6615
0.56201 | -81.5395
0.56568 | -81.6399
0.56936 | -82.0099
0.57303 | -82.0548
0.5767 | -82.1673
0.58038 | -82.1638
0.58405 | -81.7372
0.58772 | -81.6962
0.5914 | -82.1924
0.59507 | -82.4341
0.59874 | -82.4449
0.60242 | -82.8858
0.60609 | -83.6529
0.60976 | -83.8553
0.61343 | -83.3552
0.61711 | -82.8373
0.62078 | -82.754
0.62445 | -82.9187
0.62813 | -83.0681
0.6318 | -83.4664
0.63548 | -83.5761
0.63915 | -83.2729
0.64282 | -82.7473
0.64649 | -82.5406
0.65017 | -83.1559
0.65384 | -84.1425
0.65751 | -84.684
0.66119 | -84.8192
0.66486 | -84.4963
0.66853 | -84.0025
0.67221 | -83.8356
0.67588 | -84.1552
0.67955 | -84.7072
0.68323 | -84.6202
0.6869 | -84.0565
0.69057 | -83.7098
0.69425 | -84.1254
0.69792 | -85.0143
0.70159 | -85.3897
0.70527 | -85.1003
0.70894 | -84.4449
0.71261 | -83.9147
0.71629 | -83.9518
0.71996 | -84.2349
0.72363 | -83.9037
0.72731 | -83.437
0.73098 | -83.3223
0.73465 | -83.3736
0.73833 | -83.5618
0.742 | -83.7681
0.74567 | -83.9351
0.74935 | -84.0571
0.75302 | -84.0489
0.75669 | -84.1369
0.76036 | -84.547
0.76404 | -85.0351
0.76771 | -85.0769
0.77138 | -84.9753
0.77506 | -84.8551
0.77873 | -84.7082
0.78241 | -85.0428
0.78608 | -85.4549
0.78975 | -85.5103
0.79343 | -85.045
0.7971 | -84.9946
0.80077 | -85.3806
0.80444 | -85.1483
0.80812 | -84.5378
0.81179 | -84.5966
0.81546 | -85.2954
0.81914 | -85.8231
0.82281 | -85.7352
0.82648 | -85.4545
0.83016 | -85.6152
0.83383 | -85.8564
0.8375 | -85.5795
0.84118 | -85.1351
0.84485 | -85.1407
0.84852 | -85.739
0.8522 | -86.4378
0.85587 | -86.631
0.85954 | -86.2871
0.86322 | -85.8301
0.86689 | -85.5205
0.87056 | -85.3887
0.87424 | -85.5748
0.87791 | -86.0028
0.88158 | -86.1822
0.88526 | -86.1376
0.88893 | -86.2126
0.8926 | -86.3246
0.89628 | -86.3613
0.89995 | -86.3693
0.90362 | -86.3788
0.9073 | -86.5408
0.91097 | -86.7998
0.91464 | -86.6004
0.91832 | -86.2717
0.92199 | -86.2154
0.92566 | -86.4576
0.92934 | -86.7391
0.93301 | -86.8948
0.93668 | -87.1772
0.94036 | -87.2826
0.94403 | -87.0374
0.9477 | -86.8267
0.95137 | -86.959
0.95505 | -87.1791
0.95872 | -87.3002
0.9624 | -87.2975
0.96607 | -87.2167
0.96974 | -87.3582
0.97342 | -87.0486
0.97709 | -86.3977
0.98076 | -86.4861
0.98443 | -87.2469
0.98811 | -87.6294
0.99178 | -87.4973
0.99545 | -86.9999
0.99913 | -86.2804
1.0028 | -86.2351
1.0065 | -86.5746
1.0101 | -86.9847
1.0138 | -87.3032
1.0175 | -87.6544
1.0212 | -87.9622
1.0248 | -88.1243
1.0285 | -88.1311
1.0322 | -88.1248
1.0359 | -87.9572
1.0395 | -87.7067
1.0432 | -87.6673
1.0469 | -87.8979
1.0506 | -87.9365
1.0542 | -87.8498
1.0579 | -87.9559
1.0616 | -87.9679
1.0652 | -87.8714
1.0689 | -87.6472
1.0726 | -87.4014
1.0763 | -87.3742
1.0799 | -87.5664
1.0836 | -87.7335
1.0873 | -87.833
1.091 | -87.9356
1.0946 | -88.036
1.0983 | -88.1877
1.102 | -88.3
1.1057 | -88.2349
1.1093 | -88.1037
1.113 | -87.9856
1.1167 | -87.8731
1.1203 | -87.883
1.124 | -88.0481
1.1277 | -88.2223
1.1314 | -88.3092
1.135 | -88.3374
1.1387 | -88.4194
1.1424 | -88.6269
1.1461 | -88.8748
1.1497 | -89.1059
1.1534 | -89.3029
1.1571 | -89.2676
1.1608 | -88.9449
1.1644 | -88.6182
1.1681 | -88.4952
1.1718 | -88.5029
1.1754 | -88.4847
1.1791 | -88.4961
1.1828 | -88.6774
1.1865 | -88.9974
1.1901 | -89.3584
1.1938 | -89.6206
1.1975 | -89.5788
1.2012 | -89.2007
1.2048 | -88.7551
1.2085 | -88.5909
1.2122 | -88.8714
1.2159 | -89.4803
1.2195 | -90.0119
1.2232 | -90.2142
1.2269 | -90.2226
1.2305 | -90.0527
1.2342 | -89.7512
1.2379 | -89.4964
1.2416 | -89.3217
1.2452 | -89.2341
1.2489 | -89.2411
1.2526 | -89.2954
1.2563 | -89.3656
1.2599 | -89.4205
1.2636 | -89.3916
1.2673 | -89.3359
1.2709 | -89.455
1.2746 | -89.78
1.2783 | -90.0961
1.282 | -90.1048
1.2856 | -89.6079
1.2893 | -88.942
1.293 | -88.6682
1.2967 | -88.9157
1.3003 | -89.4171
1.304 | -89.7676
1.3077 | -89.7078
1.3114 | -89.4111
1.315 | -89.2765
1.3187 | -89.4948
1.3224 | -89.9061
1.326 | -90.1314
1.3297 | -90.0331
1.3334 | -89.8248
1.3371 | -89.7236
1.3407 | -89.752
1.3444 | -89.7998
1.3481 | -89.8538
1.3518 | -89.9734
1.3554 | -90.1459
1.3591 | -90.4353
1.3628 | -90.8724
1.3665 | -91.1876
1.3701 | -91.1105
1.3738 | -90.7827
1.3775 | -90.5819
1.3811 | -90.6732
1.3848 | -90.8949
1.3885 | -91.0647
1.3922 | -91.1887
1.3958 | -91.3108
1.3995 | -91.3038
1.4032 | -90.9787
1.4069 | -90.5626
1.4105 | -90.4336
1.4142 | -90.5533
1.4179 | -90.6053
1.4216 | -90.4682
1.4252 | -90.3447
1.4289 | -90.4632
1.4326 | -90.8476
1.4362 | -91.2595
1.4399 | -91.4089
1.4436 | -91.345
1.4473 | -91.2803
1.4509 | -91.2176
1.4546 | -91.1601
1.4583 | -91.188
1.462 | -91.2342
1.4656 | -91.178
1.4693 | -91.0458
1.473 | -91.0137
1.4767 | -91.1471
1.4803 | -91.263
1.484 | -91.2582
1.4877 | -91.1808
1.4913 | -90.956
1.495 | -90.7357
1.4987 | -90.8306
1.5024 | -91.1654
1.506 | -91.4572
1.5097 | -91.7095
1.5134 | -91.9242
1.5171 | -91.9252
1.5207 | -91.7137
1.5244 | -91.4552
1.5281 | -91.2365
1.5318 | -91.0839
1.5354 | -91.0353
1.5391 | -91.1668
1.5428 | -91.4373
1.5464 | -91.5737
1.5501 | -91.3382
1.5538 | -90.9005
1.5575 | -90.5648
1.5611 | -90.4316
1.5648 | -90.4885
1.5685 | -90.7537
1.5722 | -91.2489
1.5758 | -91.7422
1.5795 | -91.7369
1.5832 | -91.3226
1.5869 | -91.2412
1.5905 | -91.2377
1.5942 | -91.3025
1.5979 | -91.3692
1.6015 | -91.4298
1.6052 | -91.4928
1.6089 | -91.559
1.6126 | -91.631
1.6162 | -91.6832
1.6199 | -91.6822
1.6236 | -91.6699
1.6273 | -91.6795
1.6309 | -91.6982
1.6346 | -91.7291
1.6383 | -91.7963
1.6419 | -91.8841
1.6456 | -91.9952
1.6493 | -92.1733
1.653 | -92.4024
1.6566 | -92.5473
1.6603 | -92.4705
1.664 | -92.1682
1.6677 | -91.7765
1.6713 | -91.4508
1.675 | -91.3069
1.6787 | -91.3432
1.6824 | -91.5089
1.686 | -91.7076
1.6897 | -91.8759
1.6934 | -91.9722
1.697 | -91.9842
1.7007 | -91.9463
1.7044 | -91.9276
1.7081 | -91.9367
1.7117 | -91.9701
1.7154 | -92.017
1.7191 | -92.0849
1.7228 | -92.1444
1.7264 | -92.1454
1.7301 | -92.0371
1.7338 | -91.8476
1.7375 | -91.7052
1.7411 | -91.7533
1.7448 | -92.0009
1.7485 | -92.3112
1.7521 | -92.4974
1.7558 | -92.494
1.7595 | -92.3357
1.7632 | -92.0624
1.7668 | -91.7701
1.7705 | -91.5802
1.7742 | -91.5632
1.7779 | -91.6799
1.7815 | -91.8539
1.7852 | -92.0411
1.7889 | -92.2478
1.7926 | -92.5003
1.7962 | -92.7572
1.7999 | -92.888
1.8036 | -92.8073
1.8072 | -92.5607
1.8109 | -92.3413
1.8146 | -92.2807
1.8183 | -92.3823
1.8219 | -92.5123
1.8256 | -92.5898
1.8293 | -92.6381
1.833 | -92.6962
1.8366 | -92.7285
1.8403 | -92.6928
1.844 | -92.5975
1.8477 | -92.5092
1.8513 | -92.4448
1.855 | -92.3761
1.8587 | -92.2713
1.8623 | -92.2081
1.866 | -92.3149
1.8697 | -92.633
1.8734 | -92.9927
1.877 | -93.1244
1.8807 | -92.9494
1.8844 | -92.6778
1.8881 | -92.5343
1.8917 | -92.5844
1.8954 | -92.7582
1.8991 | -92.9337
1.9027 | -93.0494
1.9064 | -93.1135
1.9101 | -93.1476
1.9138 | -93.1859
1.9174 | -93.2416
1.9211 | -93.2838
1.9248 | -93.2667
1.9285 | -93.1962
1.9321 | -93.1271
1.9358 | -93.0969
1.9395 | -93.1364
1.9432 | -93.2597
1.9468 | -93.472
1.9505 | -93.7166
1.9542 | -93.9078
1.9578 | -94.0423
1.9615 | -94.1829
1.9652 | -94.2833
1.9689 | -94.2057
1.9725 | -93.9328
1.9762 | -93.6099
1.9799 | -93.3788
1.9836 | -93.2669
1.9872 | -93.2095
1.9909 | -93.1469
1.9946 | -93.1051
1.9983 | -93.1632
2.0019 | -93.2976
2.0056 | -93.3912
2.0093 | -93.4143
2.0129 | -93.4445
2.0166 | -93.5478
2.0203 | -93.6799
2.024 | -93.7755
2.0276 | -93.8118
2.0313 | -93.8256
2.035 | -93.8352
2.0387 | -93.7996
2.0423 | -93.686
2.046 | -93.5504
2.0497 | -93.4661
2.0534 | -93.4565
2.057 | -93.4716
2.0607 | -93.4264
2.0644 | -93.2747
2.068 | -93.0538
2.0717 | -92.8692
2.0754 | -92.8824
2.0791 | -93.1682
2.0827 | -93.6202
2.0864 | -93.9559
2.0901 | -94.005
2.0938 | -93.9127
2.0974 | -93.8369
2.1011 | -93.793
2.1048 | -93.7087
2.1085 | -93.5395
2.1121 | -93.3227
2.1158 | -93.0811
2.1195 | -92.8748
2.1231 | -92.7318
2.1268 | -92.7053
2.1305 | -92.8127
2.1342 | -93.0424
2.1378 | -93.3552
2.1415 | -93.6919
2.1452 | -93.9691
2.1489 | -94.1271
2.1525 | -94.1463
2.1562 | -94.0423
2.1599 | -93.8747
2.1636 | -93.7269
2.1672 | -93.6507
2.1709 | -93.6715
2.1746 | -93.7697
2.1782 | -93.8769
2.1819 | -93.9581
2.1856 | -94.0193
2.1893 | -94.0926
2.1929 | -94.1946
2.1966 | -94.3005
2.2003 | -94.3322
2.204 | -94.251
2.2076 | -94.0852
2.2113 | -93.9076
2.215 | -93.7866
2.2187 | -93.7798
2.2223 | -93.8975
2.226 | -94.0996
2.2297 | -94.2818
2.2333 | -94.3578
2.237 | -94.2833
2.2407 | -94.1198
2.2444 | -93.9602
2.248 | -93.8559
2.2517 | -93.8067
2.2554 | -93.7919
2.2591 | -93.7885
2.2627 | -93.7965
2.2664 | -93.8087
2.2701 | -93.8651
2.2738 | -93.9898
2.2774 | -94.1868
2.2811 | -94.3982
2.2848 | -94.5703
2.2884 | -94.6485
2.2921 | -94.6057
2.2958 | -94.4664
2.2995 | -94.2984
2.3031 | -94.1868
2.3068 | -94.1914
2.3105 | -94.2782
2.3142 | -94.4041
2.3178 | -94.496
2.3215 | -94.4797
2.3252 | -94.3327
2.3288 | -94.1499
2.3325 | -94.0403
2.3362 | -94.0569
2.3399 | -94.1933
2.3435 | -94.3778
2.3472 | -94.5224
2.3509 | -94.615
2.3546 | -94.6702
2.3582 | -94.7109
2.3619 | -94.7324
2.3656 | -94.6944
2.3693 | -94.5698
2.3729 | -94.3973
2.3766 | -94.2607
2.3803 | -94.2151
2.3839 | -94.2535
2.3876 | -94.3262
2.3913 | -94.3741
2.395 | -94.4261
2.3986 | -94.5133
2.4023 | -94.6453
2.406 | -94.7733
2.4097 | -94.8675
2.4133 | -94.9082
2.417 | -94.8824
2.4207 | -94.7839
2.4244 | -94.6334
2.428 | -94.5046
2.4317 | -94.4668
2.4354 | -94.5091
2.439 | -94.5833
2.4427 | -94.6034
2.4464 | -94.5556
2.4501 | -94.4541
2.4537 | -94.3544
2.4574 | -94.2918
2.4611 | -94.2798
2.4648 | -94.3062
2.4684 | -94.344
2.4721 | -94.3867
2.4758 | -94.4605
2.4795 | -94.551
2.4831 | -94.6546
2.4868 | -94.773
2.4905 | -94.9269
2.4941 | -95.1265
2.4978 | -95.3378
2.5015 | -95.5033
2.5052 | -95.5638
2.5088 | -95.4977
2.5125 | -95.3538
2.5162 | -95.203
2.5199 | -95.1242
2.5235 | -95.0937
2.5272 | -95.0537
2.5309 | -94.9765
2.5346 | -94.8995
2.5382 | -94.8934
2.5419 | -94.9855
2.5456 | -95.1335
2.5492 | -95.2508
2.5529 | -95.2941
2.5566 | -95.2797
2.5603 | -95.2344
2.5639 | -95.2158
2.5676 | -95.2533
2.5713 | -95.3621
2.575 | -95.5054
2.5786 | -95.6335
2.5823 | -95.7148
2.586 | -95.7568
2.5897 | -95.7909
2.5933 | -95.8125
2.597 | -95.8111
2.6007 | -95.7612
2.6043 | -95.6591
2.608 | -95.5344
2.6117 | -95.4157
2.6154 | -95.3412
2.619 | -95.3445
2.6227 | -95.4236
2.6264 | -95.5254
2.6301 | -95.5772
2.6337 | -95.5331
2.6374 | -95.4216
2.6411 | -95.323
2.6448 | -95.3034
2.6484 | -95.4119
2.6521 | -95.4671
2.6558 | -95.5243
2.6594 | -95.5645
2.6631 | -95.5737
2.6668 | -95.5722
2.6705 | -95.585
2.6741 | -95.6322
2.6778 | -95.7407
2.6815 | -95.9009
2.6852 | -96.0772
2.6888 | -96.2017
2.6925 | -96.254
2.6962 | -96.2414
2.6998 | -96.1783
2.7035 | -96.0949
2.7072 | -96.0173
2.7109 | -95.9481
2.7145 | -95.8755
2.7182 | -95.7965
2.7219 | -95.7159
2.7256 | -95.649
2.7292 | -95.6216
2.7329 | -95.629
2.7366 | -95.6555
2.7403 | -95.6815
2.7439 | -95.7188
2.7476 | -95.7873
2.7513 | -95.886
2.7549 | -96.0053
2.7586 | -96.1262
2.7623 | -96.2232
2.766 | -96.2742
2.7696 | -96.2714
2.7733 | -96.2117
2.777 | -96.1052
2.7807 | -95.9905
2.7843 | -95.9035
2.788 | -95.8632
2.7917 | -95.8528
2.7954 | -95.8649
2.799 | -95.8935
2.8027 | -95.933
2.8064 | -95.9707
2.81 | -95.9955
2.8137 | -96.0148
2.8174 | -96.036
2.8211 | -96.0712
2.8247 | -96.1094
2.8284 | -96.1519
2.8321 | -96.1908
2.8358 | -96.2097
2.8394 | -96.2129
2.8431 | -96.1957
2.8468 | -96.1844
2.8505 | -96.1793
2.8541 | -96.1919
2.8578 | -96.2124
2.8615 | -96.223
2.8651 | -96.2369
2.8688 | -96.274
2.8725 | -96.3352
2.8762 | -96.3951
2.8798 | -96.4386
2.8835 | -96.4598
2.8872 | -96.4658
2.8909 | -96.4732
2.8945 | -96.4813
2.8982 | -96.5056
2.9019 | -96.5385
2.9056 | -96.5652
2.9092 | -96.5839
2.9129 | -96.58
2.9166 | -96.56
2.9202 | -96.5401
2.9239 | -96.5219
2.9276 | -96.5156
2.9313 | -96.5302
2.9349 | -96.5477
2.9386 | -96.5624
2.9423 | -96.5629
2.946 | -96.5348
2.9496 | -96.4837
2.9533 | -96.4359
2.957 | -96.4177
2.9606 | -96.4308
2.9643 | -96.4684
2.968 | -96.5263
2.9717 | -96.5904
2.9753 | -96.6338
2.979 | -96.6266
2.9827 | -96.5668
2.9864 | -96.4373
2.99 | -96.2526
2.9937 | -96.0662
2.9974 | -95.9354
3.0011 | -95.9097
3.0047 | -95.9833
3.0084 | -96.1307
3.0121 | -96.2745
3.0157 | -96.3618
3.0194 | -96.3819
3.0231 | -96.3675
3.0268 | -96.355
3.0304 | -96.3577
3.0341 | -96.3894
3.0378 | -96.4401
3.0415 | -96.4821
3.0451 | -96.5071
3.0488 | -96.5139
3.0525 | -96.5329
3.0562 | -96.5697
3.0598 | -96.6284
3.0635 | -96.6945
3.0672 | -96.7637
3.0708 | -96.8219
3.0745 | -96.8552
3.0782 | -96.8574
3.0819 | -96.8303
3.0855 | -96.7732
3.0892 | -96.7003
3.0929 | -96.6185
3.0966 | -96.5479
3.1002 | -96.5042
3.1039 | -96.5139
3.1076 | -96.5855
3.1113 | -96.7039
3.1149 | -96.8156
3.1186 | -96.8691
3.1223 | -96.829
3.1259 | -96.7095
3.1296 | -96.5771
3.1333 | -96.5056
3.137 | -96.509
3.1406 | -96.5901
3.1443 | -96.7214
3.148 | -96.8863
3.1517 | -97.0779
3.1553 | -97.2827
3.159 | -97.4746
3.1627 | -97.6225
3.1664 | -97.69
3.17 | -97.6937
3.1737 | -97.6571
3.1774 | -97.5763
3.181 | -97.5149
3.1847 | -97.4565
3.1884 | -97.4063
3.1921 | -97.3525
3.1957 | -97.3014
3.1994 | -97.2486
3.2031 | -97.1908
3.2068 | -97.141
3.2104 | -97.0948
3.2141 | -97.0684
3.2178 | -97.0675
3.2215 | -97.0813
3.2251 | -97.106
3.2288 | -97.1465
3.2325 | -97.1837
3.2361 | -97.2226
3.2398 | -97.2668
3.2435 | -97.3147
3.2472 | -97.3561
3.2508 | -97.3794
3.2545 | -97.39
3.2582 | -97.3931
3.2619 | -97.3756
3.2655 | -97.3558
3.2692 | -97.3219
3.2729 | -97.273
3.2766 | -97.2122
3.2802 | -97.1516
3.2839 | -97.1142
3.2876 | -97.1223
3.2912 | -97.1698
3.2949 | -97.2517
3.2986 | -97.3588
3.3023 | -97.4626
3.3059 | -97.5477
3.3096 | -97.5945
3.3133 | -97.6162
3.317 | -97.645
3.3206 | -97.6708
3.3243 | -97.6914
3.328 | -97.7108
3.3316 | -97.7265
3.3353 | -97.7197
3.339 | -97.6629
3.3427 | -97.5708
3.3463 | -97.45
3.35 | -97.3195
3.3537 | -97.207
3.3574 | -97.1279
3.361 | -97.0901
3.3647 | -97.0956
3.3684 | -97.1358
3.3721 | -97.1967
3.3757 | -97.2662
3.3794 | -97.3216
3.3831 | -97.3539
3.3867 | -97.3673
3.3904 | -97.3768
3.3941 | -97.3971
3.3978 | -97.4191
3.4014 | -97.4546
3.4051 | -97.4939
3.4088 | -97.5349
3.4125 | -97.5807
3.4161 | -97.6222
3.4198 | -97.6581
3.4235 | -97.6777
3.4272 | -97.682
3.4308 | -97.6802
3.4345 | -97.6807
3.4382 | -97.6783
3.4418 | -97.661
3.4455 | -97.6365
3.4492 | -97.5946
3.4529 | -97.5361
3.4565 | -97.464
3.4602 | -97.3984
3.4639 | -97.3457
3.4676 | -97.3101
3.4712 | -97.2984
3.4749 | -97.2978
3.4786 | -97.3108
3.4823 | -97.3365
3.4859 | -97.3632
3.4896 | -97.3937
3.4933 | -97.4252
3.4969 | -97.4655
3.5006 | -97.513
3.5043 | -97.5658
3.508 | -97.6181
3.5116 | -97.6581
3.5153 | -97.6778
3.519 | -97.6819
3.5227 | -97.6726
3.5263 | -97.6594
3.53 | -97.6431
3.5337 | -97.5985
3.5374 | -97.5463
3.541 | -97.505
3.5447 | -97.4871
3.5484 | -97.5188
3.552 | -97.5884
3.5557 | -97.684
3.5594 | -97.7956
3.5631 | -97.8794
3.5667 | -97.9232
3.5704 | -97.9179
3.5741 | -97.8793
3.5778 | -97.8391
3.5814 | -97.8134
3.5851 | -97.831
3.5888 | -97.8937
3.5925 | -97.9743
3.5961 | -98.0373
3.5998 | -98.0518
3.6035 | -97.9978
3.6071 | -97.8921
3.6108 | -97.7557
3.6145 | -97.6136
3.6182 | -97.5166
3.6218 | -97.4558
3.6255 | -97.4347
3.6292 | -97.4513
3.6329 | -97.4978
3.6365 | -97.5707
3.6402 | -97.674
3.6439 | -97.8121
3.6475 | -97.9644
3.6512 | -98.1207
3.6549 | -98.2628
3.6586 | -98.3834
3.6622 | -98.4757
3.6659 | -98.5416
3.6696 | -98.575
3.6733 | -98.5594
3.6769 | -98.5028
3.6806 | -98.396
3.6843 | -98.2384
3.688 | -98.0432
3.6916 | -97.8527
3.6953 | -97.689
3.699 | -97.5766
3.7026 | -97.5273
3.7063 | -97.5901
3.71 | -97.6272
3.7137 | -97.6929
3.7173 | -97.7733
3.721 | -97.8719
3.7247 | -97.9713
3.7284 | -98.0697
3.732 | -98.1465
3.7357 | -98.1921
3.7394 | -98.2099
3.7431 | -98.2085
3.7467 | -98.2011
3.7504 | -98.1857
3.7541 | -98.1664
3.7577 | -98.1429
3.7614 | -98.1162
3.7651 | -98.097
3.7688 | -98.0951
3.7724 | -98.1258
3.7761 | -98.2042
3.7798 | -98.3222
3.7835 | -98.4565
3.7871 | -98.5957
3.7908 | -98.7241
3.7945 | -98.8248
3.7982 | -98.8938
3.8018 | -98.9322
3.8055 | -98.9523
3.8092 | -98.9655
3.8128 | -98.9648
3.8165 | -98.9592
3.8202 | -98.9521
3.8239 | -98.9491
3.8275 | -98.942
3.8312 | -98.9299
3.8349 | -98.8961
3.8386 | -98.8301
3.8422 | -98.7355
3.8459 | -98.6151
3.8496 | -98.4847
3.8533 | -98.344
3.8569 | -98.2287
3.8606 | -98.1453
3.8643 | -98.0916
3.8679 | -98.074
3.8716 | -98.0901
3.8753 | -98.1352
3.879 | -98.1985
3.8826 | -98.2721
3.8863 | -98.335
3.89 | -98.3847
3.8937 | -98.4256
3.8973 | -98.4552
3.901 | -98.4618
3.9047 | -98.4618
3.9084 | -98.4518
3.912 | -98.4281
3.9157 | -98.4036
3.9194 | -98.3696
3.923 | -98.3346
3.9267 | -98.3088
3.9304 | -98.2817
3.9341 | -98.2564
3.9377 | -98.227
3.9414 | -98.2005
3.9451 | -98.1808
3.9488 | -98.1695
3.9524 | -98.1744
3.9561 | -98.1857
3.9598 | -98.2129
3.9635 | -98.2642
3.9671 | -98.3204
3.9708 | -98.3902
3.9745 | -98.4753
3.9781 | -98.5705
3.9818 | -98.665
3.9855 | -98.7424
3.9892 | -98.8038
3.9928 | -98.86
3.9965 | -98.9048
4.0002 | -98.948
4.0039 | -98.9876
4.0075 | -99.0228
4.0112 | -99.061
4.0149 | -99.0857
4.0186 | -99.0981
4.0222 | -99.0833
4.0259 | -99.0432
4.0296 | -98.9876
4.0332 | -98.9103
4.0369 | -98.8188
4.0406 | -98.7306
4.0443 | -98.6567
4.0479 | -98.6154
4.0516 | -98.6013
4.0553 | -98.618
4.059 | -98.6684
4.0626 | -98.7331
4.0663 | -98.7987
4.07 | -98.8625
4.0736 | -98.9056
4.0773 | -98.924
4.081 | -98.9182
4.0847 | -98.8864
4.0883 | -98.8155
4.092 | -98.7268
4.0957 | -98.6352
4.0994 | -98.551
4.103 | -98.476
4.1067 | -98.4317
4.1104 | -98.422
4.1141 | -98.4524
4.1177 | -98.5227
4.1214 | -98.6347
4.1251 | -98.7795
4.1287 | -98.9301
4.1324 | -99.0689
4.1361 | -99.1869
4.1398 | -99.2503
4.1434 | -99.255
4.1471 | -99.2043
4.1508 | -99.109
4.1545 | -98.9687
4.1581 | -98.8142
4.1618 | -98.6625
4.1655 | -98.5344
4.1692 | -98.4482
4.1728 | -98.4114
4.1765 | -98.4208
4.1802 | -98.4617
4.1838 | -98.518
4.1875 | -98.582
4.1912 | -98.6379
4.1949 | -98.686
4.1985 | -98.7239
4.2022 | -98.761
4.2059 | -98.7865
4.2096 | -98.7959
4.2132 | -98.791
4.2169 | -98.7887
4.2206 | -98.7932
4.2243 | -98.8144
4.2279 | -98.8465
4.2316 | -98.8806
4.2353 | -98.9083
4.2389 | -98.9427
4.2426 | -98.976
4.2463 | -99.0093
4.25 | -99.0522
4.2536 | -99.0859
4.2573 | -99.1174
4.261 | -99.1301
4.2647 | -99.1183
4.2683 | -99.0859
4.272 | -99.0336
4.2757 | -98.9675
4.2794 | -98.9028
4.283 | -98.8317
4.2867 | -98.7685
4.2904 | -98.7203
4.294 | -98.6954
4.2977 | -98.6894
4.3014 | -98.7068
4.3051 | -98.7395
4.3087 | -98.7883
4.3124 | -98.8452
4.3161 | -98.9055
4.3198 | -98.9574
4.3234 | -99.011
4.3271 | -99.0613
4.3308 | -99.1061
4.3345 | -99.1575
4.3381 | -99.2124
4.3418 | -99.2763
4.3455 | -99.3509
4.3491 | -99.4243
4.3528 | -99.5016
4.3565 | -99.5701
4.3602 | -99.6245
4.3638 | -99.6603
4.3675 | -99.6704
4.3712 | -99.6454
4.3749 | -99.5878
4.3785 | -99.4948
4.3822 | -99.3831
4.3859 | -99.2642
4.3895 | -99.1523
4.3932 | -99.0569
4.3969 | -98.985
4.4006 | -98.9443
4.4042 | -98.9361
4.4079 | -98.9589
4.4116 | -99.0077
4.4153 | -99.0724
4.4189 | -99.1502
4.4226 | -99.2336
4.4263 | -99.3136
4.43 | -99.3894
4.4336 | -99.4589
4.4373 | -99.5241
4.441 | -99.5904
4.4446 | -99.6447
4.4483 | -99.6898
4.452 | -99.7251
4.4557 | -99.7432
4.4593 | -99.7404
4.463 | -99.7222
4.4667 | -99.6959
4.4704 | -99.67
4.474 | -99.6404
4.4777 | -99.6138
4.4814 | -99.5921
4.4851 | -99.5824
4.4887 | -99.5898
4.4924 | -99.6132
4.4961 | -99.6572
4.4997 | -99.7156
4.5034 | -99.7819
4.5071 | -99.8514
4.5108 | -99.9093
4.5144 | -99.9489
4.5181 | -99.9624
4.5218 | -99.943
4.5255 | -99.9024
4.5291 | -99.8569
4.5328 | -99.8109
4.5365 | -99.7697
4.5402 | -99.7412
4.5438 | -99.7283
4.5475 | -99.7283
4.5512 | -99.7354
4.5548 | -99.7348
4.5585 | -99.7147
4.5622 | -99.683
4.5659 | -99.6429
4.5695 | -99.5928
4.5732 | -99.5383
4.5769 | -99.4811
4.5806 | -99.4291
4.5842 | -99.3863
4.5879 | -99.3565
4.5916 | -99.3485
4.5953 | -99.3575
4.5989 | -99.3801
4.6026 | -99.4219
4.6063 | -99.4768
4.6099 | -99.5454
4.6136 | -99.6253
4.6173 | -99.7083
4.621 | -99.7917
4.6246 | -99.868
4.6283 | -99.9336
4.632 | -99.9702
4.6357 | -99.9808
4.6393 | -99.9718
4.643 | -99.9471
4.6467 | -99.9208
4.6504 | -99.8969
4.654 | -99.8732
4.6577 | -99.8448
4.6614 | -99.8126
4.665 | -99.7765
4.6687 | -99.7424
4.6724 | -99.7175
4.6761 | -99.7051
4.6797 | -99.6905
4.6834 | -99.6875
4.6871 | -99.702
4.6908 | -99.7201
4.6944 | -99.7447
4.6981 | -99.7683
4.7018 | -99.7859
4.7054 | -99.7822
4.7091 | -99.7695
4.7128 | -99.7471
4.7165 | -99.7103
4.7201 | -99.6775
4.7238 | -99.6639
4.7275 | -99.6691
4.7312 | -99.6906
4.7348 | -99.7317
4.7385 | -99.7826
4.7422 | -99.8379
4.7459 | -99.9021
4.7495 | -99.9692
4.7532 | -100.0201
4.7569 | -100.0463
4.7605 | -100.0542
4.7642 | -100.0349
4.7679 | -99.9914
4.7716 | -99.9681
4.7752 | -99.9513
4.7789 | -99.9384
4.7826 | -99.9294
4.7863 | -99.9348
4.7899 | -99.9564
4.7936 | -99.9876
4.7973 | -100.0305
4.801 | -100.087
4.8046 | -100.1505
4.8083 | -100.2195
4.812 | -100.2909
4.8156 | -100.3602
4.8193 | -100.4111
4.823 | -100.4554
4.8267 | -100.4837
4.8303 | -100.4859
4.834 | -100.4668
4.8377 | -100.4331
4.8414 | -100.3813
4.845 | -100.3165
4.8487 | -100.2444
4.8524 | -100.172
4.8561 | -100.1109
4.8597 | -100.0572
4.8634 | -100.01
4.8671 | -99.972
4.8707 | -99.9403
4.8744 | -99.9176
4.8781 | -99.9049
4.8818 | -99.9065
4.8854 | -99.9225
4.8891 | -99.9585
4.8928 | -100.0008
4.8965 | -100.0502
4.9001 | -100.1037
4.9038 | -100.1528
4.9075 | -100.2071
4.9112 | -100.2638
4.9148 | -100.3172
4.9185 | -100.372
4.9222 | -100.4325
4.9258 | -100.4922
4.9295 | -100.5493
4.9332 | -100.6022
4.9369 | -100.6403
4.9405 | -100.6664
4.9442 | -100.6804
4.9479 | -100.683
4.9516 | -100.6706
4.9552 | -100.6503
4.9589 | -100.623
4.9626 | -100.5837
4.9663 | -100.5482
4.9699 | -100.5243
4.9736 | -100.5052
4.9773 | -100.4881
4.9809 | -100.4785
4.9846 | -100.4778
4.9883 | -100.4834
4.992 | -100.5075
4.9956 | -100.5419
4.9993 | -100.5864
5.003 | -100.6466
5.0067 | -100.713
5.0103 | -100.7706
5.014 | -100.8093
5.0177 | -100.823
5.0214 | -100.8002
5.025 | -100.7521
5.0287 | -100.6865
5.0324 | -100.6145
5.036 | -100.5406
5.0397 | -100.483
5.0434 | -100.4324
5.0471 | -100.401
5.0507 | -100.3942
5.0544 | -100.4041
5.0581 | -100.4288
5.0618 | -100.4662
5.0654 | -100.5143
5.0691 | -100.5672
5.0728 | -100.6294
5.0765 | -100.6898
5.0801 | -100.7371
5.0838 | -100.78
5.0875 | -100.8154
5.0911 | -100.839
5.0948 | -100.8445
5.0985 | -100.8347
5.1022 | -100.8027
5.1058 | -100.7515
5.1095 | -100.6829
5.1132 | -100.6107
5.1169 | -100.5441
5.1205 | -100.4803
5.1242 | -100.4241
5.1279 | -100.3724
5.1315 | -100.3271
5.1352 | -100.2779
5.1389 | -100.2282
5.1426 | -100.1732
5.1462 | -100.1107
5.1499 | -100.0524
5.1536 | -100.0087
5.1573 | -99.9828
5.1609 | -99.9735
5.1646 | -99.99
5.1683 | -100.036
5.172 | -100.0941
5.1756 | -100.166
5.1793 | -100.2474
5.183 | -100.3354
5.1866 | -100.4176
5.1903 | -100.4973
5.194 | -100.5583
5.1977 | -100.6029
5.2013 | -100.6353
5.205 | -100.6563
5.2087 | -100.6613
5.2124 | -100.6551
5.216 | -100.6335
5.2197 | -100.6178
5.2234 | -100.6129
5.2271 | -100.5994
5.2307 | -100.5892
5.2344 | -100.5902
5.2381 | -100.5929
5.2417 | -100.5992
5.2454 | -100.6102
5.2491 | -100.6307
5.2528 | -100.6515
5.2564 | -100.665
5.2601 | -100.6817
5.2638 | -100.7033
5.2675 | -100.7237
5.2711 | -100.7502
5.2748 | -100.7828
5.2785 | -100.8097
5.2822 | -100.8355
5.2858 | -100.8616
5.2895 | -100.8769
5.2932 | -100.8849
5.2968 | -100.8826
5.3005 | -100.8852
5.3042 | -100.8799
5.3079 | -100.8772
5.3115 | -100.8718
5.3152 | -100.8687
5.3189 | -100.8672
5.3226 | -100.8637
5.3262 | -100.8724
5.3299 | -100.899
5.3336 | -100.9396
5.3373 | -100.9916
5.3409 | -101.0473
5.3446 | -101.1124
5.3483 | -101.1772
5.3519 | -101.2405
5.3556 | -101.2927
5.3593 | -101.3336
5.363 | -101.3642
5.3666 | -101.388
5.3703 | -101.4133
5.374 | -101.4262
5.3777 | -101.4299
5.3813 | -101.436
5.385 | -101.4506
5.3887 | -101.4716
5.3924 | -101.4948
5.396 | -101.5188
5.3997 | -101.5355
5.4034 | -101.5483
5.407 | -101.5497
5.4107 | -101.5332
5.4144 | -101.5009
5.4181 | -101.4719
5.4217 | -101.4481
5.4254 | -101.4235
5.4291 | -101.3965
5.4328 | -101.3645
5.4364 | -101.3255
5.4401 | -101.2902
5.4438 | -101.2503
5.4474 | -101.2032
5.4511 | -101.1542
5.4548 | -101.1023
5.4585 | -101.0561
5.4621 | -101.0166
5.4658 | -100.983
5.4695 | -100.9505
5.4732 | -100.928
5.4768 | -100.9112
5.4805 | -100.8953
5.4842 | -100.8826
5.4879 | -100.8764
5.4915 | -100.8731
5.4952 | -100.872
5.4989 | -100.8778
5.5025 | -100.8823
5.5062 | -100.8978
5.5099 | -100.9255
5.5136 | -100.9648
5.5172 | -101.0168
5.5209 | -101.0725
5.5246 | -101.1216
5.5283 | -101.1684
5.5319 | -101.2061
5.5356 | -101.2318
5.5393 | -101.2494
5.543 | -101.2537
5.5466 | -101.244
5.5503 | -101.2288
5.554 | -101.2101
5.5576 | -101.1832
5.5613 | -101.1619
5.565 | -101.1449
5.5687 | -101.1377
5.5723 | -101.1313
5.576 | -101.1308
5.5797 | -101.1333
5.5834 | -101.1358
5.587 | -101.1324
5.5907 | -101.1181
5.5944 | -101.0869
5.5981 | -101.0488
5.6017 | -101.0068
5.6054 | -100.9577
5.6091 | -100.8954
5.6127 | -100.8296
5.6164 | -100.7575
5.6201 | -100.6751
5.6238 | -100.5862
5.6274 | -100.4962
5.6311 | -100.4131
5.6348 | -100.3362
5.6385 | -100.2694
5.6421 | -100.2299
5.6458 | -100.2149
5.6495 | -100.2247
5.6532 | -100.2595
5.6568 | -100.3151
5.6605 | -100.382
5.6642 | -100.4625
5.6678 | -100.5475
5.6715 | -100.6379
5.6752 | -100.7204
5.6789 | -100.792
5.6825 | -100.8454
5.6862 | -100.8831
5.6899 | -100.9083
5.6936 | -100.9244
5.6972 | -100.9255
5.7009 | -100.9248
5.7046 | -100.927
5.7083 | -100.9299
5.7119 | -100.9334
5.7156 | -100.9373
5.7193 | -100.9407
5.7229 | -100.9473
5.7266 | -100.9448
5.7303 | -100.9401
5.734 | -100.9463
5.7376 | -100.9617
5.7413 | -100.984
5.745 | -101.0146
5.7487 | -101.0428
5.7523 | -101.0765
5.756 | -101.1084
5.7597 | -101.1463
5.7633 | -101.1913
5.767 | -101.2342
5.7707 | -101.2836
5.7744 | -101.3463
5.778 | -101.4156
5.7817 | -101.4852
5.7854 | -101.5559
5.7891 | -101.6322
5.7927 | -101.7009
5.7964 | -101.758
5.8001 | -101.7934
5.8038 | -101.8098
5.8074 | -101.8113
5.8111 | -101.7935
5.8148 | -101.7667
5.8184 | -101.7367
5.8221 | -101.7035
5.8258 | -101.7011
5.8295 | -101.697
5.8331 | -101.6972
5.8368 | -101.701
5.8405 | -101.6975
5.8442 | -101.6922
5.8478 | -101.6771
5.8515 | -101.652
5.8552 | -101.6186
5.8589 | -101.5772
5.8625 | -101.5293
5.8662 | -101.4777
5.8699 | -101.424
5.8735 | -101.3718
5.8772 | -101.3176
5.8809 | -101.271
5.8846 | -101.2268
5.8882 | -101.1915
5.8919 | -101.1683
5.8956 | -101.1479
5.8993 | -101.14
5.9029 | -101.136
5.9066 | -101.1407
5.9103 | -101.1516
5.914 | -101.1684
5.9176 | -101.1877
5.9213 | -101.2063
5.925 | -101.2274
5.9286 | -101.246
5.9323 | -101.2661
5.936 | -101.283
5.9397 | -101.3004
5.9433 | -101.3179
5.947 | -101.3376
5.9507 | -101.3596
5.9544 | -101.3942
5.958 | -101.4382
5.9617 | -101.4823
5.9654 | -101.5301
5.9691 | -101.5804
5.9727 | -101.6261
5.9764 | -101.6755
5.9801 | -101.7146
5.9837 | -101.7377
5.9874 | -101.752
5.9911 | -101.756
5.9948 | -101.7483
5.9984 | -101.7319
6.0021 | -101.7148
6.0058 | -101.6914
6.0095 | -101.6651
6.0131 | -101.6423
6.0168 | -101.62
6.0205 | -101.6038
6.0242 | -101.5911
6.0278 | -101.5756
6.0315 | -101.5582
6.0352 | -101.5354
6.0388 | -101.5116
6.0425 | -101.4803
6.0462 | -101.4422
6.0499 | -101.4056
6.0535 | -101.3639
6.0572 | -101.3136
6.0609 | -101.2645
6.0646 | -101.2252
6.0682 | -101.1879
6.0719 | -101.1658
6.0756 | -101.1554
6.0793 | -101.1566
6.0829 | -101.17
6.0866 | -101.1956
6.0903 | -101.2275
6.0939 | -101.2629
6.0976 | -101.2948
6.1013 | -101.3288
6.105 | -101.3599
6.1086 | -101.388
6.1123 | -101.4097
6.116 | -101.4247
6.1197 | -101.4354
6.1233 | -101.4429
6.127 | -101.4482
6.1307 | -101.4542
6.1344 | -101.4654
6.138 | -101.48
6.1417 | -101.5007
6.1454 | -101.5371
6.149 | -101.5765
6.1527 | -101.6171
6.1564 | -101.6545
6.1601 | -101.695
6.1637 | -101.7266
6.1674 | -101.7513
6.1711 | -101.7734
6.1748 | -101.7902
6.1784 | -101.8052
6.1821 | -101.8103
6.1858 | -101.8141
6.1894 | -101.8234
6.1931 | -101.8269
6.1968 | -101.8289
6.2005 | -101.8331
6.2041 | -101.8351
6.2078 | -101.8436
6.2115 | -101.8566
6.2152 | -101.867
6.2188 | -101.8817
6.2225 | -101.8956
6.2262 | -101.9208
6.2299 | -101.9495
6.2335 | -101.977
6.2372 | -102.0121
6.2409 | -102.0484
6.2445 | -102.0859
6.2482 | -102.1249
6.2519 | -102.1569
6.2556 | -102.1923
6.2592 | -102.2314
6.2629 | -102.263
6.2666 | -102.296
6.2703 | -102.331
6.2739 | -102.3661
6.2776 | -102.3967
6.2813 | -102.424
6.285 | -102.4414
6.2886 | -102.4419
6.2923 | -102.4293
6.296 | -102.4113
6.2996 | -102.3821
6.3033 | -102.3431
6.307 | -102.2997
6.3107 | -102.2496
6.3143 | -102.1868
6.318 | -102.1206
6.3217 | -102.0626
6.3254 | -102.0151
6.329 | -101.9716
6.3327 | -101.9401
6.3364 | -101.9209
6.3401 | -101.9191
6.3437 | -101.9284
6.3474 | -101.9479
6.3511 | -101.9741
6.3547 | -102.0058
6.3584 | -102.0395
6.3621 | -102.0658
6.3658 | -102.0913
6.3694 | -102.1153
6.3731 | -102.1353
6.3768 | -102.1533
6.3805 | -102.1709
6.3841 | -102.1961
6.3878 | -102.2197
6.3915 | -102.2368
6.3952 | -102.258
6.3988 | -102.2869
6.4025 | -102.3216
6.4062 | -102.3599
6.4098 | -102.4043
6.4135 | -102.4413
6.4172 | -102.473
6.4209 | -102.5014
6.4245 | -102.5194
6.4282 | -102.5218
6.4319 | -102.505
6.4356 | -102.4686
6.4392 | -102.4185
6.4429 | -102.3526
6.4466 | -102.2748
6.4502 | -102.1892
6.4539 | -102.1004
6.4576 | -102.0177
6.4613 | -101.9457
6.4649 | -101.8828
6.4686 | -101.8374
6.4723 | -101.8066
6.476 | -101.791
6.4796 | -101.7971
6.4833 | -101.8146
6.487 | -101.8438
6.4907 | -101.8869
6.4943 | -101.94
6.498 | -101.998
6.5017 | -102.0579
6.5053 | -102.1144
6.509 | -102.1679
6.5127 | -102.2187
6.5164 | -102.2573
6.52 | -102.2802
6.5237 | -102.2916
6.5274 | -102.3044
6.5311 | -102.3073
6.5347 | -102.3069
6.5384 | -102.3047
6.5421 | -102.3074
6.5458 | -102.316
6.5494 | -102.3313
6.5531 | -102.3574
6.5568 | -102.394
6.5604 | -102.4414
6.5641 | -102.4977
6.5678 | -102.5599
6.5715 | -102.629
6.5751 | -102.6998
6.5788 | -102.7653
6.5825 | -102.8279
6.5862 | -102.8885
6.5898 | -102.9309
6.5935 | -102.9468
6.5972 | -102.9475
6.6009 | -102.9272
6.6045 | -102.8883
6.6082 | -102.8352
6.6119 | -102.7646
6.6155 | -102.6848
6.6192 | -102.6032
6.6229 | -102.5293
6.6266 | -102.4538
6.6302 | -102.3878
6.6339 | -102.3312
6.6376 | -102.2894
6.6413 | -102.2576
6.6449 | -102.2401
6.6486 | -102.2394
6.6523 | -102.2471
6.656 | -102.2597
6.6596 | -102.2843
6.6633 | -102.3147
6.667 | -102.3482
6.6706 | -102.3848
6.6743 | -102.417
6.678 | -102.4455
6.6817 | -102.468
6.6853 | -102.4815
6.689 | -102.4955
6.6927 | -102.5038
6.6964 | -102.4981
6.7 | -102.4821
6.7037 | -102.4657
6.7074 | -102.4492
6.7111 | -102.4354
6.7147 | -102.4232
6.7184 | -102.4057
6.7221 | -102.3806
6.7257 | -102.3587
6.7294 | -102.3416
6.7331 | -102.3193
6.7368 | -102.2989
6.7404 | -102.2819
6.7441 | -102.2672
6.7478 | -102.2531
6.7515 | -102.2456
6.7551 | -102.2523
6.7588 | -102.2653
6.7625 | -102.2811
6.7662 | -102.3012
6.7698 | -102.3258
6.7735 | -102.3578
6.7772 | -102.3902
6.7808 | -102.4223
6.7845 | -102.4576
6.7882 | -102.481
6.7919 | -102.4937
6.7955 | -102.5091
6.7992 | -102.518
6.8029 | -102.5092
6.8066 | -102.4849
6.8102 | -102.4485
6.8139 | -102.4049
6.8176 | -102.3583
6.8213 | -102.3136
6.8249 | -102.2647
6.8286 | -102.211
6.8323 | -102.172
6.8359 | -102.141
6.8396 | -102.1155
6.8433 | -102.0995
6.847 | -102.0942
6.8506 | -102.097
6.8543 | -102.1069
6.858 | -102.1264
6.8617 | -102.1545
6.8653 | -102.1845
6.869 | -102.2185
6.8727 | -102.2545
6.8763 | -102.2904
6.88 | -102.3271
6.8837 | -102.3619
6.8874 | -102.397
6.891 | -102.4356
6.8947 | -102.4779
6.8984 | -102.5064
6.9021 | -102.5271
6.9057 | -102.5522
6.9094 | -102.5722
6.9131 | -102.5898
6.9168 | -102.6041
6.9204 | -102.6127
6.9241 | -102.6101
6.9278 | -102.6024
6.9314 | -102.5965
6.9351 | -102.5813
6.9388 | -102.5624
6.9425 | -102.5404
6.9461 | -102.5128
6.9498 | -102.4807
6.9535 | -102.4479
6.9572 | -102.4122
6.9608 | -102.3763
6.9645 | -102.3437
6.9682 | -102.3097
6.9719 | -102.2736
6.9755 | -102.2394
6.9792 | -102.2134
6.9829 | -102.194
6.9865 | -102.1888
6.9902 | -102.1959
6.9939 | -102.2149
6.9976 | -102.2387
7.0012 | -102.2703
7.0049 | -102.316
7.0086 | -102.3677
7.0123 | -102.4155
7.0159 | -102.4625
7.0196 | -102.5014
7.0233 | -102.5305
7.027 | -102.5522
7.0306 | -102.5664
7.0343 | -102.5707
7.038 | -102.5634
7.0416 | -102.5469
7.0453 | -102.528
7.049 | -102.5117
7.0527 | -102.4957
7.0563 | -102.48
7.06 | -102.4646
7.0637 | -102.4556
7.0674 | -102.4563
7.071 | -102.4656
7.0747 | -102.4826
7.0784 | -102.5068
7.0821 | -102.5382
7.0857 | -102.5718
7.0894 | -102.6042
7.0931 | -102.6391
7.0967 | -102.6774
7.1004 | -102.7147
7.1041 | -102.7541
7.1078 | -102.7897
7.1114 | -102.8199
7.1151 | -102.8461
7.1188 | -102.87
7.1225 | -102.888
7.1261 | -102.9061
7.1298 | -102.9195
7.1335 | -102.9196
7.1372 | -102.9164
7.1408 | -102.9108
7.1445 | -102.8964
7.1482 | -102.8779
7.1518 | -102.8586
7.1555 | -102.8345
7.1592 | -102.8086
7.1629 | -102.7753
7.1665 | -102.7407
7.1702 | -102.7056
7.1739 | -102.663
7.1776 | -102.6248
7.1812 | -102.5906
7.1849 | -102.5639
7.1886 | -102.5443
7.1922 | -102.5317
7.1959 | -102.5284
7.1996 | -102.5282
7.2033 | -102.5368
7.2069 | -102.5589
7.2106 | -102.5914
7.2143 | -102.625
7.218 | -102.6599
7.2216 | -102.6955
7.2253 | -102.7329
7.229 | -102.7767
7.2327 | -102.8131
7.2363 | -102.8419
7.24 | -102.8723
7.2437 | -102.8895
7.2473 | -102.8949
7.251 | -102.8933
7.2547 | -102.8851
7.2584 | -102.87
7.262 | -102.8582
7.2657 | -102.8446
7.2694 | -102.8414
7.2731 | -102.84
7.2767 | -102.8429
7.2804 | -102.8521
7.2841 | -102.8602
7.2878 | -102.8734
7.2914 | -102.8773
7.2951 | -102.8757
7.2988 | -102.8691
7.3024 | -102.8539
7.3061 | -102.8308
7.3098 | -102.7951
7.3135 | -102.7551
7.3171 | -102.7095
7.3208 | -102.6638
7.3245 | -102.6189
7.3282 | -102.5795
7.3318 | -102.5434
7.3355 | -102.5119
7.3392 | -102.4911
7.3429 | -102.4787
7.3465 | -102.4716
7.3502 | -102.4771
7.3539 | -102.4902
7.3575 | -102.508
7.3612 | -102.5333
7.3649 | -102.5648
7.3686 | -102.6019
7.3722 | -102.6471
7.3759 | -102.6956
7.3796 | -102.7481
7.3833 | -102.7994
7.3869 | -102.8506
7.3906 | -102.8974
7.3943 | -102.9337
7.398 | -102.9659
7.4016 | -102.99
7.4053 | -103.0048
7.409 | -103.0105
7.4126 | -103.0114
7.4163 | -102.9936
7.42 | -102.9885
7.4237 | -102.9765
7.4273 | -102.9683
7.431 | -102.9653
7.4347 | -102.9625
7.4384 | -102.9634
7.442 | -102.9682
7.4457 | -102.9753
7.4494 | -102.9885
7.4531 | -103.0044
7.4567 | -103.0238
7.4604 | -103.0403
7.4641 | -103.058
7.4677 | -103.0841
7.4714 | -103.107
7.4751 | -103.126
7.4788 | -103.1461
7.4824 | -103.1627
7.4861 | -103.1767
7.4898 | -103.1866
7.4935 | -103.1944
7.4971 | -103.1977
7.5008 | -103.1974
7.5045 | -103.1928
7.5081 | -103.1878
7.5118 | -103.1806
7.5155 | -103.1734
7.5192 | -103.1668
7.5228 | -103.1559
7.5265 | -103.1484
7.5302 | -103.1478
7.5339 | -103.1414
7.5375 | -103.1359
7.5412 | -103.1247
7.5449 | -103.1078
7.5486 | -103.0963
7.5522 | -103.0789
7.5559 | -103.0611
7.5596 | -103.0452
7.5632 | -103.0249
7.5669 | -103.0028
7.5706 | -102.9772
7.5743 | -102.9572
7.5779 | -102.9384
7.5816 | -102.9166
7.5853 | -102.905
7.589 | -102.8994
7.5926 | -102.8983
7.5963 | -102.9075
7.6 | -102.921
7.6037 | -102.9417
7.6073 | -102.964
7.611 | -102.9944
7.6147 | -103.0264
7.6183 | -103.0584
7.622 | -103.0875
7.6257 | -103.112
7.6294 | -103.1321
7.633 | -103.1477
7.6367 | -103.1601
7.6404 | -103.1678
7.6441 | -103.1653
7.6477 | -103.162
7.6514 | -103.1568
7.6551 | -103.1495
7.6588 | -103.1422
7.6624 | -103.1258
7.6661 | -103.1111
7.6698 | -103.0971
7.6734 | -103.0851
7.6771 | -103.079
7.6808 | -103.0813
7.6845 | -103.0867
7.6881 | -103.0892
7.6918 | -103.095
7.6955 | -103.0986
7.6992 | -103.1013
7.7028 | -103.1058
7.7065 | -103.1061
7.7102 | -103.0994
7.7139 | -103.0911
7.7175 | -103.0787
7.7212 | -103.0665
7.7249 | -103.0509
7.7285 | -103.0325
7.7322 | -103.0202
7.7359 | -103.0093
7.7396 | -103.0041
7.7432 | -103.0083
7.7469 | -103.0129
7.7506 | -103.0201
7.7543 | -103.0298
7.7579 | -103.0448
7.7616 | -103.0631
7.7653 | -103.0777
7.769 | -103.086
7.7726 | -103.0996
7.7763 | -103.1164
7.78 | -103.1254
7.7836 | -103.1349
7.7873 | -103.1451
7.791 | -103.1488
7.7947 | -103.1451
7.7983 | -103.1465
7.802 | -103.1494
7.8057 | -103.1517
7.8094 | -103.1599
7.813 | -103.1641
7.8167 | -103.1647
7.8204 | -103.1692
7.8241 | -103.1774
7.8277 | -103.1864
7.8314 | -103.1953
7.8351 | -103.1981
7.8387 | -103.1953
7.8424 | -103.1949
7.8461 | -103.187
7.8498 | -103.1768
7.8534 | -103.1672
7.8571 | -103.1526
7.8608 | -103.1377
7.8645 | -103.1225
7.8681 | -103.1104
7.8718 | -103.0974
7.8755 | -103.085
7.8792 | -103.0824
7.8828 | -103.084
7.8865 | -103.0799
7.8902 | -103.0779
7.8938 | -103.0864
7.8975 | -103.0951
7.9012 | -103.1033
7.9049 | -103.1214
7.9085 | -103.1406
7.9122 | -103.1554
7.9159 | -103.1738
7.9196 | -103.1914
7.9232 | -103.2044
7.9269 | -103.2238
7.9306 | -103.2477
7.9342 | -103.2671
7.9379 | -103.2838
7.9416 | -103.2968
7.9453 | -103.3118
7.9489 | -103.3182
7.9526 | -103.3324
7.9563 | -103.347
7.96 | -103.3581
7.9636 | -103.3684
7.9673 | -103.3782
7.971 | -103.3866
7.9747 | -103.3943
7.9783 | -103.4108
7.982 | -103.4235
7.9857 | -103.4269
7.9893 | -103.4337
7.993 | -103.4425
7.9967 | -103.443
8.0004 | -103.4458
8.004 | -103.4485
8.0077 | -103.4478
8.0114 | -103.4419
8.0151 | -103.4355
8.0187 | -103.4293
8.0224 | -103.4158
8.0261 | -103.4034
8.0298 | -103.3934
8.0334 | -103.3858
8.0371 | -103.3803
8.0408 | -103.3828
8.0444 | -103.3856
8.0481 | -103.3889
8.0518 | -103.3963
8.0555 | -103.4071
8.0591 | -103.4241
8.0628 | -103.4404
8.0665 | -103.4587
8.0702 | -103.4697
8.0738 | -103.4795
8.0775 | -103.4918
8.0812 | -103.5033
8.0849 | -103.5144
8.0885 | -103.527
8.0922 | -103.54
8.0959 | -103.557
8.0995 | -103.5738
8.1032 | -103.592
8.1069 | -103.6127
8.1106 | -103.6296
8.1142 | -103.6407
8.1179 | -103.6487
8.1216 | -103.652
8.1253 | -103.6507
8.1289 | -103.6523
8.1326 | -103.6545
8.1363 | -103.6552
8.14 | -103.6553
8.1436 | -103.6593
8.1473 | -103.6569
8.151 | -103.6551
8.1546 | -103.6537
8.1583 | -103.6511
8.162 | -103.6441
8.1657 | -103.6365
8.1693 | -103.6318
8.173 | -103.6222
8.1767 | -103.6052
8.1804 | -103.5792
8.184 | -103.5525
8.1877 | -103.5283
8.1914 | -103.5016
8.1951 | -103.4688
8.1987 | -103.4313
8.2024 | -103.3936
8.2061 | -103.3596
8.2097 | -103.3274
8.2134 | -103.2995
8.2171 | -103.2774
8.2208 | -103.2579
8.2244 | -103.2482
8.2281 | -103.2445
8.2318 | -103.2445
8.2355 | -103.2529
8.2391 | -103.2706
8.2428 | -103.2957
8.2465 | -103.3278
8.2501 | -103.36
8.2538 | -103.3993
8.2575 | -103.4456
8.2612 | -103.4933
8.2648 | -103.5382
8.2685 | -103.5854
8.2722 | -103.627
8.2759 | -103.6638
8.2795 | -103.6962
8.2832 | -103.7208
8.2869 | -103.7362
8.2906 | -103.7495
8.2942 | -103.7608
8.2979 | -103.7676
8.3016 | -103.7755
8.3052 | -103.7836
8.3089 | -103.7904
8.3126 | -103.806
8.3163 | -103.8291
8.3199 | -103.8517
8.3236 | -103.8781
8.3273 | -103.9049
8.331 | -103.9343
8.3346 | -103.9641
8.3383 | -103.9893
8.342 | -104.0112
8.3457 | -104.0298
8.3493 | -104.0483
8.353 | -104.0567
8.3567 | -104.0508
8.3603 | -104.0382
8.364 | -104.0134
8.3677 | -103.9787
8.3714 | -103.9427
8.375 | -103.9015
8.3787 | -103.855
8.3824 | -103.8045
8.3861 | -103.754
8.3897 | -103.7077
8.3934 | -103.6599
8.3971 | -103.6209
8.4008 | -103.5898
8.4044 | -103.5676
8.4081 | -103.5537
8.4118 | -103.5458
8.4154 | -103.544
8.4191 | -103.5468
8.4228 | -103.5487
8.4265 | -103.5514
8.4301 | -103.5572
8.4338 | -103.5663
8.4375 | -103.5756
8.4412 | -103.5839
8.4448 | -103.5922
8.4485 | -103.5994
8.4522 | -103.6048
8.4559 | -103.6058
8.4595 | -103.6063
8.4632 | -103.5993
8.4669 | -103.5842
8.4705 | -103.5606
8.4742 | -103.539
8.4779 | -103.5145
8.4816 | -103.4984
8.4852 | -103.4798
8.4889 | -103.4624
8.4926 | -103.4488
8.4963 | -103.4404
8.4999 | -103.4374
8.5036 | -103.4416
8.5073 | -103.4592
8.511 | -103.485
8.5146 | -103.5131
8.5183 | -103.5487
8.522 | -103.5856
8.5256 | -103.6294
8.5293 | -103.6719
8.533 | -103.7121
8.5367 | -103.748
8.5403 | -103.7859
8.544 | -103.8208
8.5477 | -103.8504
8.5514 | -103.8675
8.555 | -103.8734
8.5587 | -103.8657
8.5624 | -103.8538
8.5661 | -103.8368
8.5697 | -103.8122
8.5734 | -103.781
8.5771 | -103.741
8.5807 | -103.7021
8.5844 | -103.665
8.5881 | -103.6339
8.5918 | -103.6023
8.5954 | -103.5747
8.5991 | -103.5498
8.6028 | -103.533
8.6065 | -103.5222
8.6101 | -103.5157
8.6138 | -103.5195
8.6175 | -103.5312
8.6211 | -103.546
8.6248 | -103.5611
8.6285 | -103.5799
8.6322 | -103.6046
8.6358 | -103.6302
8.6395 | -103.652
8.6432 | -103.674
8.6469 | -103.7001
8.6505 | -103.7262
8.6542 | -103.7506
8.6579 | -103.777
8.6616 | -103.8032
8.6652 | -103.8247
8.6689 | -103.8404
8.6726 | -103.8515
8.6762 | -103.8592
8.6799 | -103.8697
8.6836 | -103.8747
8.6873 | -103.8754
8.6909 | -103.8694
8.6946 | -103.8612
8.6983 | -103.8491
8.702 | -103.8327
8.7056 | -103.8219
8.7093 | -103.8085
8.713 | -103.7917
8.7167 | -103.7762
8.7203 | -103.7703
8.724 | -103.7662
8.7277 | -103.7676
8.7313 | -103.7699
8.735 | -103.7811
8.7387 | -103.7975
8.7424 | -103.8166
8.746 | -103.8366
8.7497 | -103.863
8.7534 | -103.8915
8.7571 | -103.9177
8.7607 | -103.9368
8.7644 | -103.9498
8.7681 | -103.9662
8.7718 | -103.9805
8.7754 | -103.9828
8.7791 | -103.9788
8.7828 | -103.9694
8.7864 | -103.9551
8.7901 | -103.9378
8.7938 | -103.9172
8.7975 | -103.8947
8.8011 | -103.8759
8.8048 | -103.8562
8.8085 | -103.8368
8.8122 | -103.8168
8.8158 | -103.8005
8.8195 | -103.7896
8.8232 | -103.7822
8.8269 | -103.7789
8.8305 | -103.7822
8.8342 | -103.782
8.8379 | -103.7845
8.8415 | -103.7903
8.8452 | -103.7966
8.8489 | -103.7987
8.8526 | -103.8045
8.8562 | -103.8109
8.8599 | -103.8185
8.8636 | -103.8267
8.8673 | -103.8385
8.8709 | -103.8522
8.8746 | -103.869
8.8783 | -103.8853
8.882 | -103.9095
8.8856 | -103.9367
8.8893 | -103.9629
8.893 | -103.9979
8.8966 | -104.0266
8.9003 | -104.0478
8.904 | -104.0715
8.9077 | -104.0922
8.9113 | -104.1139
8.915 | -104.1289
8.9187 | -104.1376
8.9224 | -104.1375
8.926 | -104.1283
8.9297 | -104.1105
8.9334 | -104.0859
8.937 | -104.0447
8.9407 | -104.0071
8.9444 | -103.9657
8.9481 | -103.9161
8.9517 | -103.8658
8.9554 | -103.8132
8.9591 | -103.7623
8.9628 | -103.7179
8.9664 | -103.6767
8.9701 | -103.6337
8.9738 | -103.5987
8.9775 | -103.5667
8.9811 | -103.5384
8.9848 | -103.5151
8.9885 | -103.4966
8.9921 | -103.4828
8.9958 | -103.4719
8.9995 | -103.4655
9.0032 | -103.4677
9.0068 | -103.4879
9.0105 | -103.5005
9.0142 | -103.5146
9.0179 | -103.535
9.0215 | -103.5649
9.0252 | -103.5905
9.0289 | -103.6187
9.0326 | -103.6528
9.0362 | -103.6878
9.0399 | -103.7235
9.0436 | -103.7544
9.0472 | -103.7844
9.0509 | -103.8149
9.0546 | -103.846
9.0583 | -103.8799
9.0619 | -103.9084
9.0656 | -103.9363
9.0693 | -103.9711
9.073 | -104.0022
9.0766 | -104.0284
9.0803 | -104.056
9.084 | -104.0806
9.0877 | -104.1027
9.0913 | -104.1195
9.095 | -104.1358
9.0987 | -104.1493
9.1023 | -104.1612
9.106 | -104.1645
9.1097 | -104.1616
9.1134 | -104.1552
9.117 | -104.149
9.1207 | -104.1425
9.1244 | -104.13
9.1281 | -104.1204
9.1317 | -104.1123
9.1354 | -104.1034
9.1391 | -104.095
9.1428 | -104.0878
9.1464 | -104.0816
9.1501 | -104.0779
9.1538 | -104.0757
9.1574 | -104.0741
9.1611 | -104.0775
9.1648 | -104.0813
9.1685 | -104.0801
9.1721 | -104.0786
9.1758 | -104.0761
9.1795 | -104.0723
9.1832 | -104.0721
9.1868 | -104.0708
9.1905 | -104.0665
9.1942 | -104.0524
9.1979 | -104.0381
9.2015 | -104.0205
9.2052 | -104.0008
9.2089 | -103.9749
9.2125 | -103.9525
9.2162 | -103.9254
9.2199 | -103.8978
9.2236 | -103.8741
9.2272 | -103.8487
9.2309 | -103.8199
9.2346 | -103.7948
9.2383 | -103.7753
9.2419 | -103.7575
9.2456 | -103.7437
9.2493 | -103.733
9.253 | -103.7245
9.2566 | -103.7214
9.2603 | -103.7268
9.264 | -103.7367
9.2676 | -103.7479
9.2713 | -103.7613
9.275 | -103.7757
9.2787 | -103.7933
9.2823 | -103.8134
9.286 | -103.8349
9.2897 | -103.8568
9.2934 | -103.8755
9.297 | -103.893
9.3007 | -103.909
9.3044 | -103.9225
9.308 | -103.9273
9.3117 | -103.9292
9.3154 | -103.9261
9.3191 | -103.9206
9.3227 | -103.9137
9.3264 | -103.9055
9.3301 | -103.9038
9.3338 | -103.9051
9.3374 | -103.9053
9.3411 | -103.9076
9.3448 | -103.9136
9.3485 | -103.9248
9.3521 | -103.9454
9.3558 | -103.9677
9.3595 | -103.9899
9.3631 | -104.0188
9.3668 | -104.0534
9.3705 | -104.0904
9.3742 | -104.1289
9.3778 | -104.169
9.3815 | -104.212
9.3852 | -104.2561
9.3889 | -104.2966
9.3925 | -104.3365
9.3962 | -104.376
9.3999 | -104.4112
9.4036 | -104.449
9.4072 | -104.4867
9.4109 | -104.5201
9.4146 | -104.5524
9.4182 | -104.5817
9.4219 | -104.6039
9.4256 | -104.6198
9.4293 | -104.6328
9.4329 | -104.6417
9.4366 | -104.6496
9.4403 | -104.6521
9.444 | -104.6501
9.4476 | -104.6377
9.4513 | -104.6179
9.455 | -104.5904
9.4587 | -104.5604
9.4623 | -104.5286
9.466 | -104.488
9.4697 | -104.444
9.4733 | -104.4025
9.477 | -104.3593
9.4807 | -104.3165
9.4844 | -104.2687
9.488 | -104.2257
9.4917 | -104.1913
9.4954 | -104.1612
9.4991 | -104.1327
9.5027 | -104.1028
9.5064 | -104.0754
9.5101 | -104.0481
9.5138 | -104.0249
9.5174 | -104.0074
9.5211 | -103.9918
9.5248 | -103.9817
9.5284 | -103.9754
9.5321 | -103.9768
9.5358 | -104.001
9.5395 | -104.0124
9.5431 | -104.0261
9.5468 | -104.0418
9.5505 | -104.0618
9.5542 | -104.0777
9.5578 | -104.0962
9.5615 | -104.1188
9.5652 | -104.1431
9.5689 | -104.1624
9.5725 | -104.1751
9.5762 | -104.1834
9.5799 | -104.1903
9.5835 | -104.1934
9.5872 | -104.1902
9.5909 | -104.1881
9.5946 | -104.1798
9.5982 | -104.1658
9.6019 | -104.1445
9.6056 | -104.1201
9.6093 | -104.0931
9.6129 | -104.0602
9.6166 | -104.0303
9.6203 | -104.0031
9.624 | -103.9778
9.6276 | -103.9512
9.6313 | -103.9267
9.635 | -103.9055
9.6386 | -103.8939
9.6423 | -103.8893
9.646 | -103.8916
9.6497 | -103.9008
9.6533 | -103.9159
9.657 | -103.9333
9.6607 | -103.9546
9.6644 | -103.9819
9.668 | -104.0109
9.6717 | -104.0377
9.6754 | -104.0616
9.679 | -104.0851
9.6827 | -104.1104
9.6864 | -104.1341
9.6901 | -104.1601
9.6937 | -104.18
9.6974 | -104.1992
9.7011 | -104.2114
9.7048 | -104.2198
9.7084 | -104.2281
9.7121 | -104.2337
9.7158 | -104.2329
9.7195 | -104.2303
9.7231 | -104.2231
9.7268 | -104.214
9.7305 | -104.1991
9.7341 | -104.1767
9.7378 | -104.1568
9.7415 | -104.1344
9.7452 | -104.1157
9.7488 | -104.0932
9.7525 | -104.0687
9.7562 | -104.0471
9.7599 | -104.0216
9.7635 | -103.9996
9.7672 | -103.9804
9.7709 | -103.9659
9.7746 | -103.9554
9.7782 | -103.9498
9.7819 | -103.9441
9.7856 | -103.9423
9.7892 | -103.9461
9.7929 | -103.9559
9.7966 | -103.9642
9.8003 | -103.9736
9.8039 | -103.9866
9.8076 | -104.0034
9.8113 | -104.0239
9.815 | -104.0445
9.8186 | -104.0647
9.8223 | -104.0821
9.826 | -104.0909
9.8297 | -104.0959
9.8333 | -104.1024
9.837 | -104.1056
9.8407 | -104.1024
9.8443 | -104.1016
9.848 | -104.0992
9.8517 | -104.0907
9.8554 | -104.0848
9.859 | -104.0776
9.8627 | -104.0673
9.8664 | -104.0598
9.8701 | -104.0479
9.8737 | -104.0386
9.8774 | -104.0295
9.8811 | -104.0153
9.8848 | -104.0042
9.8884 | -103.9959
9.8921 | -103.9953
9.8958 | -103.9978
9.8994 | -104.0018
9.9031 | -104.0078
9.9068 | -104.0165
9.9105 | -104.0297
9.9141 | -104.0474
9.9178 | -104.0645
9.9215 | -104.082
9.9252 | -104.1045
9.9288 | -104.1308
9.9325 | -104.1535
9.9362 | -104.1743
9.9398 | -104.1995
9.9435 | -104.2266
9.9472 | -104.2482
9.9509 | -104.2649
9.9545 | -104.2799
9.9582 | -104.295
9.9619 | -104.3076
9.9656 | -104.3207
9.9692 | -104.3354
9.9729 | -104.3436
9.9766 | -104.3398
9.9803 | -104.3399
9.9839 | -104.3405
9.9876 | -104.3358
9.9913 | -104.3286
9.9949 | -104.3249
9.9986 | -104.3175
10.0023 | -104.3107
10.006 | -104.3037
10.0096 | -104.2927
10.0133 | -104.2895
10.017 | -104.2886
10.0207 | -104.285
10.0243 | -104.2866
10.028 | -104.2918
10.0317 | -104.2962
10.0354 | -104.2989
10.039 | -104.3061
10.0427 | -104.3198
10.0464 | -104.3364
10.05 | -104.3516
10.0537 | -104.3704
10.0574 | -104.3854
10.0611 | -104.4013
10.0647 | -104.4187
10.0684 | -104.4296
10.0721 | -104.4425
10.0758 | -104.4501
10.0794 | -104.459
10.0831 | -104.4696
10.0868 | -104.474
10.0905 | -104.4741
10.0941 | -104.4777
10.0978 | -104.482
10.1015 | -104.4843
10.1051 | -104.4882
10.1088 | -104.4901
10.1125 | -104.4875
10.1162 | -104.4854
10.1198 | -104.4861
10.1235 | -104.487
10.1272 | -104.4844
10.1309 | -104.479
10.1345 | -104.4722
10.1382 | -104.4684
10.1419 | -104.4629
10.1456 | -104.4598
10.1492 | -104.4573
10.1529 | -104.4501
10.1566 | -104.4426
10.1602 | -104.433
10.1639 | -104.4204
10.1676 | -104.4086
10.1713 | -104.3953
10.1749 | -104.3806
10.1786 | -104.3689
10.1823 | -104.3572
10.186 | -104.3435
10.1896 | -104.3299
10.1933 | -104.3162
10.197 | -104.308
10.2007 | -104.2999
10.2043 | -104.2968
10.208 | -104.2972
10.2117 | -104.2989
10.2153 | -104.3006
10.219 | -104.3015
10.2227 | -104.3036
10.2264 | -104.3107
10.23 | -104.3171
10.2337 | -104.3259
10.2374 | -104.34
10.2411 | -104.3538
10.2447 | -104.3666
10.2484 | -104.3749
10.2521 | -104.3847
10.2558 | -104.3944
10.2594 | -104.4076
10.2631 | -104.4191
10.2668 | -104.4307
10.2704 | -104.4438
10.2741 | -104.4584
10.2778 | -104.4689
10.2815 | -104.4769
10.2851 | -104.4821
10.2888 | -104.4878
10.2925 | -104.4944
10.2962 | -104.4978
10.2998 | -104.5005
10.3035 | -104.5045
10.3072 | -104.5091
10.3109 | -104.5114
10.3145 | -104.5197
10.3182 | -104.5256
10.3219 | -104.528
10.3255 | -104.5265
10.3292 | -104.5251
10.3329 | -104.5221
10.3366 | -104.5184
10.3402 | -104.5136
10.3439 | -104.5099
10.3476 | -104.5052
10.3513 | -104.5009
10.3549 | -104.4978
10.3586 | -104.4974
10.3623 | -104.497
10.3659 | -104.498
10.3696 | -104.5014
10.3733 | -104.5073
10.377 | -104.5158
10.3806 | -104.5296
10.3843 | -104.5399
10.388 | -104.5516
10.3917 | -104.5678
10.3953 | -104.5831
10.399 | -104.6012
10.4027 | -104.6147
10.4064 | -104.6264
10.41 | -104.6364
10.4137 | -104.6442
10.4174 | -104.6452
10.421 | -104.6412
10.4247 | -104.6402
10.4284 | -104.6453
10.4321 | -104.6473
10.4357 | -104.6467
10.4394 | -104.6438
10.4431 | -104.6386
10.4468 | -104.6368
10.4504 | -104.6325
10.4541 | -104.6284
10.4578 | -104.6245
10.4615 | -104.6228
10.4651 | -104.6149
10.4688 | -104.611
10.4725 | -104.6081
10.4761 | -104.6059
10.4798 | -104.6045
10.4835 | -104.6024
10.4872 | -104.5981
10.4908 | -104.5889
10.4945 | -104.5831
10.4982 | -104.5756
10.5019 | -104.5647
10.5055 | -104.5544
10.5092 | -104.5443
10.5129 | -104.5354
10.5166 | -104.527
10.5202 | -104.518
10.5239 | -104.5077
10.5276 | -104.4953
10.5312 | -104.4813
10.5349 | -104.4718
10.5386 | -104.4601
10.5423 | -104.4497
10.5459 | -104.4474
10.5496 | -104.443
10.5533 | -104.4404
10.557 | -104.4392
10.5606 | -104.441
10.5643 | -104.4466
10.568 | -104.4528
10.5717 | -104.4635
10.5753 | -104.469
10.579 | -104.4725
10.5827 | -104.4828
10.5863 | -104.4968
10.59 | -104.5053
10.5937 | -104.5153
10.5974 | -104.5238
10.601 | -104.5303
10.6047 | -104.5394
10.6084 | -104.5483
10.6121 | -104.5591
10.6157 | -104.5654
10.6194 | -104.5747
10.6231 | -104.585
10.6268 | -104.5947
10.6304 | -104.6051
10.6341 | -104.6124
10.6378 | -104.6226
10.6414 | -104.6343
10.6451 | -104.6482
10.6488 | -104.6546
10.6525 | -104.661
10.6561 | -104.6687
10.6598 | -104.6716
10.6635 | -104.6705
10.6672 | -104.6673
10.6708 | -104.6605
10.6745 | -104.6524
10.6782 | -104.6476
10.6818 | -104.6371
10.6855 | -104.6254
10.6892 | -104.6137
10.6929 | -104.6024
10.6965 | -104.5896
10.7002 | -104.5735
10.7039 | -104.5629
10.7076 | -104.5515
10.7112 | -104.5384
10.7149 | -104.5233
10.7186 | -104.5101
10.7223 | -104.5024
10.7259 | -104.4882
10.7296 | -104.4762
10.7333 | -104.4669
10.7369 | -104.4579
10.7406 | -104.4514
10.7443 | -104.4493
10.748 | -104.4463
10.7516 | -104.4468
10.7553 | -104.4481
10.759 | -104.4524
10.7627 | -104.4592
10.7663 | -104.4621
10.77 | -104.4629
10.7737 | -104.4685
10.7774 | -104.4755
10.781 | -104.4778
10.7847 | -104.4814
10.7884 | -104.487
10.792 | -104.4944
10.7957 | -104.4971
10.7994 | -104.4994
10.8031 | -104.5014
10.8067 | -104.5044
10.8104 | -104.5096
10.8141 | -104.5167
10.8178 | -104.5262
10.8214 | -104.5395
10.8251 | -104.5541
10.8288 | -104.5682
10.8325 | -104.5812
10.8361 | -104.5927
10.8398 | -104.61
10.8435 | -104.6268
10.8471 | -104.6443
10.8508 | -104.6665
10.8545 | -104.6876
10.8582 | -104.709
10.8618 | -104.7313
10.8655 | -104.7547
10.8692 | -104.776
10.8729 | -104.7967
10.8765 | -104.8182
10.8802 | -104.8357
10.8839 | -104.8565
10.8876 | -104.8786
10.8912 | -104.8986
10.8949 | -104.9155
10.8986 | -104.9262
10.9022 | -104.9364
10.9059 | -104.9461
10.9096 | -104.9526
10.9133 | -104.96
10.9169 | -104.9629
10.9206 | -104.959
10.9243 | -104.9557
10.928 | -104.9503
10.9316 | -104.9391
10.9353 | -104.9238
10.939 | -104.9079
10.9427 | -104.8905
10.9463 | -104.8679
10.95 | -104.8408
10.9537 | -104.814
10.9573 | -104.7837
10.961 | -104.7533
10.9647 | -104.7236
10.9684 | -104.6914
10.972 | -104.6604
10.9757 | -104.6326
10.9794 | -104.6064
10.9831 | -104.5824
10.9867 | -104.5613
10.9904 | -104.5438
10.9941 | -104.5279
10.9977 | -104.5162
11.0014 | -104.5111
11.0051 | -104.5088
11.0088 | -104.5041
11.0124 | -104.504
11.0161 | -104.5105
11.0198 | -104.5181
11.0235 | -104.5269
11.0271 | -104.5378
11.0308 | -104.5479
11.0345 | -104.556
11.0382 | -104.5646
11.0418 | -104.5705
11.0455 | -104.5726
11.0492 | -104.5739
11.0528 | -104.5756
11.0565 | -104.5799
11.0602 | -104.5824
11.0639 | -104.5835
11.0675 | -104.5809
11.0712 | -104.576
11.0749 | -104.5723
11.0786 | -104.5678
11.0822 | -104.5669
11.0859 | -104.5643
11.0896 | -104.5637
11.0933 | -104.562
11.0969 | -104.5593
11.1006 | -104.5539
11.1043 | -104.5544
11.1079 | -104.5627
11.1116 | -104.5759
11.1153 | -104.591
11.119 | -104.6041
11.1226 | -104.6221
11.1263 | -104.6514
11.13 | -104.6717
11.1337 | -104.6931
11.1373 | -104.7137
11.141 | -104.7363
11.1447 | -104.7601
11.1484 | -104.787
11.152 | -104.8113
11.1557 | -104.8313
11.1594 | -104.8459
11.163 | -104.857
11.1667 | -104.8671
11.1704 | -104.8794
11.1741 | -104.8854
11.1777 | -104.8828
11.1814 | -104.88
11.1851 | -104.8733
11.1888 | -104.8622
11.1924 | -104.8521
11.1961 | -104.8416
11.1998 | -104.8297
11.2035 | -104.818
11.2071 | -104.8062
11.2108 | -104.7906
11.2145 | -104.7741
11.2181 | -104.7603
11.2218 | -104.7464
11.2255 | -104.7352
11.2292 | -104.7279
11.2328 | -104.7222
11.2365 | -104.7181
11.2402 | -104.7196
11.2439 | -104.7213
11.2475 | -104.7271
11.2512 | -104.7367
11.2549 | -104.7525
11.2586 | -104.7704
11.2622 | -104.7894
11.2659 | -104.8088
11.2696 | -104.8225
11.2732 | -104.835
11.2769 | -104.8502
11.2806 | -104.8688
11.2843 | -104.8845
11.2879 | -104.8956
11.2916 | -104.9048
11.2953 | -104.9124
11.299 | -104.9158
11.3026 | -104.914
11.3063 | -104.9104
11.31 | -104.9045
11.3137 | -104.8907
11.3173 | -104.8762
11.321 | -104.864
11.3247 | -104.8474
11.3283 | -104.8293
11.332 | -104.8111
11.3357 | -104.7895
11.3394 | -104.7676
11.343 | -104.7473
11.3467 | -104.7241
11.3504 | -104.7022
11.3541 | -104.6828
11.3577 | -104.6621
11.3614 | -104.6401
11.3651 | -104.6214
11.3688 | -104.6059
11.3724 | -104.5903
11.3761 | -104.5774
11.3798 | -104.5687
11.3834 | -104.5619
11.3871 | -104.559
11.3908 | -104.5571
11.3945 | -104.5576
11.3981 | -104.5618
11.4018 | -104.567
11.4055 | -104.5737
11.4092 | -104.5817
11.4128 | -104.591
11.4165 | -104.6
11.4202 | -104.6095
11.4238 | -104.6201
11.4275 | -104.6304
11.4312 | -104.6367
11.4349 | -104.6452
11.4385 | -104.6526
11.4422 | -104.6582
11.4459 | -104.6635
11.4496 | -104.6687
11.4532 | -104.6771
11.4569 | -104.6838
11.4606 | -104.6889
11.4643 | -104.6962
11.4679 | -104.7021
11.4716 | -104.7044
11.4753 | -104.7075
11.4789 | -104.7147
11.4826 | -104.7204
11.4863 | -104.7243
11.49 | -104.7321
11.4936 | -104.7349
11.4973 | -104.7349
11.501 | -104.7356
11.5047 | -104.7384
11.5083 | -104.7401
11.512 | -104.7413
11.5157 | -104.7419
11.5194 | -104.7398
11.523 | -104.7377
11.5267 | -104.7347
11.5304 | -104.7305
11.534 | -104.726
11.5377 | -104.7208
11.5414 | -104.7192
11.5451 | -104.7165
11.5487 | -104.7149
11.5524 | -104.7148
11.5561 | -104.7177
11.5598 | -104.7244
11.5634 | -104.7317
11.5671 | -104.7413
11.5708 | -104.7536
11.5745 | -104.7694
11.5781 | -104.7843
11.5818 | -104.8002
11.5855 | -104.8185
11.5891 | -104.8391
11.5928 | -104.8636
11.5965 | -104.8851
11.6002 | -104.9028
11.6038 | -104.9165
11.6075 | -104.929
11.6112 | -104.9382
11.6149 | -104.9413
11.6185 | -104.9411
11.6222 | -104.9397
11.6259 | -104.9327
11.6296 | -104.9213
11.6332 | -104.909
11.6369 | -104.8936
11.6406 | -104.869
11.6442 | -104.8391
11.6479 | -104.8051
11.6516 | -104.7664
11.6553 | -104.7194
11.6589 | -104.6828
11.6626 | -104.6467
11.6663 | -104.6118
11.67 | -104.578
11.6736 | -104.5454
11.6773 | -104.5133
11.681 | -104.4835
11.6846 | -104.4552
11.6883 | -104.4294
11.692 | -104.4049
11.6957 | -104.3879
11.6993 | -104.3748
11.703 | -104.3617
11.7067 | -104.3541
11.7104 | -104.3458
11.714 | -104.3405
11.7177 | -104.3373
11.7214 | -104.3376
11.7251 | -104.339
11.7287 | -104.3425
11.7324 | -104.349
11.7361 | -104.3558
11.7397 | -104.3628
11.7434 | -104.3723
11.7471 | -104.3866
11.7508 | -104.4029
11.7544 | -104.4154
11.7581 | -104.4285
11.7618 | -104.4427
11.7655 | -104.4598
11.7691 | -104.4803
11.7728 | -104.5018
11.7765 | -104.5185
11.7802 | -104.5347
11.7838 | -104.5547
11.7875 | -104.5753
11.7912 | -104.5972
11.7948 | -104.6205
11.7985 | -104.6425
11.8022 | -104.6655
11.8059 | -104.6875
11.8095 | -104.7136
11.8132 | -104.7411
11.8169 | -104.7675
11.8206 | -104.7956
11.8242 | -104.8228
11.8279 | -104.8456
11.8316 | -104.8658
11.8353 | -104.8886
11.8389 | -104.9083
11.8426 | -104.9256
11.8463 | -104.9386
11.8499 | -104.9507
11.8536 | -104.9589
11.8573 | -104.966
11.861 | -104.9749
11.8646 | -104.9812
11.8683 | -104.9798
11.872 | -104.976
11.8757 | -104.9765
11.8793 | -104.9728
11.883 | -104.9691
11.8867 | -104.9654
11.8904 | -104.958
11.894 | -104.951
11.8977 | -104.9449
11.9014 | -104.9383
11.905 | -104.9339
11.9087 | -104.9297
11.9124 | -104.9257
11.9161 | -104.92
11.9197 | -104.9169
11.9234 | -104.9111
11.9271 | -104.9076
11.9308 | -104.9063
11.9344 | -104.9025
11.9381 | -104.8988
11.9418 | -104.896
11.9455 | -104.8951
11.9491 | -104.8946
11.9528 | -104.8953
11.9565 | -104.896
11.9601 | -104.8961
11.9638 | -104.8961
11.9675 | -104.8998
11.9712 | -104.9037
11.9748 | -104.9035
11.9785 | -104.9027
11.9822 | -104.9031
11.9859 | -104.9041
11.9895 | -104.9062
11.9932 | -104.9044
11.9969 | -104.9042
12.0006 | -104.9049
12.0042 | -104.9044
12.0079 | -104.903
12.0116 | -104.8992
12.0152 | -104.8949
12.0189 | -104.8945
12.0226 | -104.8949
12.0263 | -104.8971
12.0299 | -104.8982
12.0336 | -104.8943
12.0373 | -104.8922
12.041 | -104.8902
12.0446 | -104.8882
12.0483 | -104.8846
12.052 | -104.8797
12.0556 | -104.8731
12.0593 | -104.867
12.063 | -104.8627
12.0667 | -104.8568
12.0703 | -104.8507
12.074 | -104.8432
12.0777 | -104.8334
12.0814 | -104.8249
12.085 | -104.817
12.0887 | -104.8068
12.0924 | -104.792
12.0961 | -104.7789
12.0997 | -104.7675
12.1034 | -104.7524
12.1071 | -104.7372
12.1107 | -104.7222
12.1144 | -104.7068
12.1181 | -104.6922
12.1218 | -104.6787
12.1254 | -104.6644
12.1291 | -104.6487
12.1328 | -104.6333
12.1365 | -104.6205
12.1401 | -104.6065
12.1438 | -104.591
12.1475 | -104.5803
12.1512 | -104.5709
12.1548 | -104.5601
12.1585 | -104.5462
12.1622 | -104.5379
12.1658 | -104.5321
12.1695 | -104.5203
12.1732 | -104.5104
12.1769 | -104.5045
12.1805 | -104.4987
12.1842 | -104.4935
12.1879 | -104.4872
12.1916 | -104.4841
12.1952 | -104.4781
12.1989 | -104.4741
12.2026 | -104.4702
12.2063 | -104.4661
12.2099 | -104.4657
12.2136 | -104.4667
12.2173 | -104.4702
12.2209 | -104.473
12.2246 | -104.4752
12.2283 | -104.4828
12.232 | -104.4913
12.2356 | -104.5
12.2393 | -104.5131
12.243 | -104.5263
12.2467 | -104.5405
12.2503 | -104.5582
12.254 | -104.5791
12.2577 | -104.6038
12.2614 | -104.632
12.265 | -104.6637
12.2687 | -104.6973
12.2724 | -104.7331
12.276 | -104.7701
12.2797 | -104.8082
12.2834 | -104.8487
12.2871 | -104.8883
12.2907 | -104.9231
12.2944 | -104.9574
12.2981 | -104.9932
12.3018 | -105.026
12.3054 | -105.0533
12.3091 | -105.0777
12.3128 | -105.0986
12.3165 | -105.1135
12.3201 | -105.1269
12.3238 | -105.1353
12.3275 | -105.1374
12.3311 | -105.1352
12.3348 | -105.1341
12.3385 | -105.1262
12.3422 | -105.1143
12.3458 | -105.1038
12.3495 | -105.0896
12.3532 | -105.0697
12.3569 | -105.0494
12.3605 | -105.0344
12.3642 | -105.0194
12.3679 | -105.0033
12.3716 | -104.9891
12.3752 | -104.977
12.3789 | -104.9602
12.3826 | -104.9484
12.3862 | -104.9395
12.3899 | -104.9324
12.3936 | -104.9271
12.3973 | -104.9231
12.4009 | -104.9238
12.4046 | -104.9263
12.4083 | -104.9286
12.412 | -104.9362
12.4156 | -104.945
12.4193 | -104.9553
12.423 | -104.9693
12.4267 | -104.9848
12.4303 | -105.0013
12.434 | -105.0177
12.4377 | -105.0335
12.4413 | -105.0485
12.445 | -105.0644
12.4487 | -105.082
12.4524 | -105.0969
12.456 | -105.112
12.4597 | -105.1264
12.4634 | -105.1363
12.4671 | -105.1441
12.4707 | -105.1536
12.4744 | -105.1644
12.4781 | -105.1725
12.4817 | -105.179
12.4854 | -105.1856
12.4891 | -105.1911
12.4928 | -105.1908
12.4964 | -105.1903
12.5001 | -105.1879
12.5038 | -105.1846
12.5075 | -105.1806
12.5111 | -105.1792
12.5148 | -105.1766
12.5185 | -105.1695
12.5222 | -105.1645
12.5258 | -105.1589
12.5295 | -105.1521
12.5332 | -105.1463
12.5368 | -105.1389
12.5405 | -105.1277
12.5442 | -105.1178
12.5479 | -105.1071
12.5515 | -105.0958
12.5552 | -105.0846
12.5589 | -105.0781
12.5626 | -105.0705
12.5662 | -105.062
12.5699 | -105.0551
12.5736 | -105.0499
12.5773 | -105.0468
12.5809 | -105.0444
12.5846 | -105.0446
12.5883 | -105.0445
12.5919 | -105.0433
12.5956 | -105.0462
12.5993 | -105.0512
12.603 | -105.0607
12.6066 | -105.0728
12.6103 | -105.0825
12.614 | -105.0919
12.6177 | -105.1042
12.6213 | -105.1186
12.625 | -105.1336
12.6287 | -105.1497
12.6324 | -105.1671
12.636 | -105.1815
12.6397 | -105.1933
12.6434 | -105.2069
12.647 | -105.219
12.6507 | -105.225
12.6544 | -105.2251
12.6581 | -105.2233
12.6617 | -105.2239
12.6654 | -105.2247
12.6691 | -105.2223
12.6728 | -105.2208
12.6764 | -105.2154
12.6801 | -105.205
12.6838 | -105.1907
12.6875 | -105.1709
12.6911 | -105.1444
12.6948 | -105.1188
12.6985 | -105.0954
12.7021 | -105.0659
12.7058 | -105.0351
12.7095 | -105.0046
12.7132 | -104.9742
12.7168 | -104.9427
12.7205 | -104.9165
12.7242 | -104.8919
12.7279 | -104.8671
12.7315 | -104.8448
12.7352 | -104.8267
12.7389 | -104.8125
12.7425 | -104.8044
12.7462 | -104.7989
12.7499 | -104.7976
12.7536 | -104.7968
12.7572 | -104.799
12.7609 | -104.8071
12.7646 | -104.8185
12.7683 | -104.8323
12.7719 | -104.8495
12.7756 | -104.8703
12.7793 | -104.8932
12.783 | -104.9147
12.7866 | -104.9373
12.7903 | -104.9594
12.794 | -104.9816
12.7976 | -105.0051
12.8013 | -105.0234
12.805 | -105.0384
12.8087 | -105.0521
12.8123 | -105.0649
12.816 | -105.0742
12.8197 | -105.0803
12.8234 | -105.0872
12.827 | -105.09
12.8307 | -105.0908
12.8344 | -105.0916
12.8381 | -105.0878
12.8417 | -105.0811
12.8454 | -105.0674
12.8491 | -105.0542
12.8527 | -105.044
12.8564 | -105.03
12.8601 | -105.0167
12.8638 | -105.0062
12.8674 | -104.9945
12.8711 | -104.9876
12.8748 | -104.9803
12.8785 | -104.9708
12.8821 | -104.9648
12.8858 | -104.9683
12.8895 | -104.9758
12.8932 | -104.9807
12.8968 | -104.9871
12.9005 | -104.9973
12.9042 | -105.0108
12.9078 | -105.0212
12.9115 | -105.0344
12.9152 | -105.0507
12.9189 | -105.0669
12.9225 | -105.0833
12.9262 | -105.0987
12.9299 | -105.1155
12.9336 | -105.1332
12.9372 | -105.151
12.9409 | -105.1702
12.9446 | -105.1853
12.9483 | -105.1977
12.9519 | -105.2119
12.9556 | -105.2253
12.9593 | -105.2373
12.9629 | -105.2482
12.9666 | -105.2555
12.9703 | -105.2625
12.974 | -105.2609
12.9776 | -105.2606
12.9813 | -105.2634
12.985 | -105.2602
12.9887 | -105.2565
12.9923 | -105.2517
12.996 | -105.2433
12.9997 | -105.232
13.0034 | -105.2199
13.007 | -105.2067
13.0107 | -105.1879
13.0144 | -105.1703
13.018 | -105.1525
13.0217 | -105.1334
13.0254 | -105.1147
13.0291 | -105.0941
13.0327 | -105.0792
13.0364 | -105.0618
13.0401 | -105.0427
13.0438 | -105.0283
13.0474 | -105.0106
13.0511 | -104.9941
13.0548 | -104.9818
13.0585 | -104.97
13.0621 | -104.9625
13.0658 | -104.9563
13.0695 | -104.9489
13.0731 | -104.9389
13.0768 | -104.932
13.0805 | -104.9315
13.0842 | -104.9316
13.0878 | -104.934
13.0915 | -104.9379
13.0952 | -104.9447
13.0989 | -104.9514
13.1025 | -104.9564
13.1062 | -104.9611
13.1099 | -104.9694
13.1136 | -104.9779
13.1172 | -104.9807
13.1209 | -104.9857
13.1246 | -104.9922
13.1282 | -104.9959
13.1319 | -104.9998
13.1356 | -105.0003
13.1393 | -104.9993
13.1429 | -104.9985
13.1466 | -104.9962
13.1503 | -104.9917
13.154 | -104.9849
13.1576 | -104.9768
13.1613 | -104.9698
13.165 | -104.9563
13.1686 | -104.9394
13.1723 | -104.9236
13.176 | -104.9036
13.1797 | -104.8801
13.1833 | -104.8577
13.187 | -104.8317
13.1907 | -104.8066
13.1944 | -104.7849
13.198 | -104.7645
13.2017 | -104.7428
13.2054 | -104.7211
13.2091 | -104.7002
13.2127 | -104.6809
13.2164 | -104.665
13.2201 | -104.6464
13.2237 | -104.631
13.2274 | -104.6176
13.2311 | -104.6027
13.2348 | -104.5969
13.2384 | -104.5916
13.2421 | -104.5875
13.2458 | -104.5949
13.2495 | -104.5922
13.2531 | -104.5904
13.2568 | -104.5885
13.2605 | -104.5866
13.2642 | -104.5873
13.2678 | -104.5903
13.2715 | -104.5933
13.2752 | -104.5974
13.2788 | -104.6022
13.2825 | -104.6056
13.2862 | -104.6072
13.2899 | -104.6098
13.2935 | -104.6152
13.2972 | -104.6197
13.3009 | -104.6226
13.3046 | -104.6263
13.3082 | -104.6304
13.3119 | -104.6329
13.3156 | -104.6333
13.3193 | -104.6348
13.3229 | -104.6378
13.3266 | -104.6369
13.3303 | -104.6318
13.3339 | -104.6276
13.3376 | -104.6263
13.3413 | -104.6245
13.345 | -104.6257
13.3486 | -104.6251
13.3523 | -104.6239
13.356 | -104.6231
13.3597 | -104.6225
13.3633 | -104.6225
13.367 | -104.6221
13.3707 | -104.6224
13.3744 | -104.6236
13.378 | -104.6297
13.3817 | -104.6338
13.3854 | -104.6337
13.389 | -104.6361
13.3927 | -104.6411
13.3964 | -104.6452
13.4001 | -104.6506
13.4037 | -104.6602
13.4074 | -104.6691
13.4111 | -104.6801
13.4148 | -104.6921
13.4184 | -104.7023
13.4221 | -104.7166
13.4258 | -104.7292
13.4294 | -104.7415
13.4331 | -104.7514
13.4368 | -104.7597
13.4405 | -104.7721
13.4441 | -104.7861
13.4478 | -104.7988
13.4515 | -104.8094
13.4552 | -104.8191
13.4588 | -104.8257
13.4625 | -104.8294
13.4662 | -104.8313
13.4699 | -104.8333
13.4735 | -104.833
13.4772 | -104.8335
13.4809 | -104.8312
13.4845 | -104.8291
13.4882 | -104.827
13.4919 | -104.8237
13.4956 | -104.817
13.4992 | -104.8095
13.5029 | -104.8043
13.5066 | -104.7983
13.5103 | -104.7928
13.5139 | -104.79
13.5176 | -104.7889
13.5213 | -104.7897
13.525 | -104.7917
13.5286 | -104.7918
13.5323 | -104.7889
13.536 | -104.7875
13.5396 | -104.7868
13.5433 | -104.7913
13.547 | -104.7982
13.5507 | -104.8053
13.5543 | -104.8103
13.558 | -104.8156
13.5617 | -104.8237
13.5654 | -104.8288
13.569 | -104.8341
13.5727 | -104.8361
13.5764 | -104.8424
13.5801 | -104.8535
13.5837 | -104.8645
13.5874 | -104.8742
13.5911 | -104.8826
13.5947 | -104.8899
13.5984 | -104.8924
13.6021 | -104.897
13.6058 | -104.9002
13.6094 | -104.9022
13.6131 | -104.9038
13.6168 | -104.9019
13.6205 | -104.8997
13.6241 | -104.8959
13.6278 | -104.8912
13.6315 | -104.8836
13.6352 | -104.8753
13.6388 | -104.866
13.6425 | -104.8564
13.6462 | -104.8491
13.6498 | -104.8411
13.6535 | -104.8319
13.6572 | -104.8179
13.6609 | -104.8051
13.6645 | -104.7953
13.6682 | -104.7862
13.6719 | -104.7784
13.6756 | -104.7742
13.6792 | -104.7679
13.6829 | -104.7632
13.6866 | -104.7626
13.6903 | -104.7628
13.6939 | -104.7634
13.6976 | -104.7649
13.7013 | -104.7701
13.7049 | -104.7801
13.7086 | -104.7885
13.7123 | -104.7983
13.716 | -104.8123
13.7196 | -104.8281
13.7233 | -104.8466
13.727 | -104.8689
13.7307 | -104.8926
13.7343 | -104.9177
13.738 | -104.941
13.7417 | -104.9654
13.7454 | -104.9917
13.749 | -105.0213
13.7527 | -105.0486
13.7564 | -105.0736
13.76 | -105.1039
13.7637 | -105.1302
13.7674 | -105.1569
13.7711 | -105.1796
13.7747 | -105.1968
13.7784 | -105.216
13.7821 | -105.2354
13.7858 | -105.2527
13.7894 | -105.2704
13.7931 | -105.2873
13.7968 | -105.302
13.8004 | -105.3119
13.8041 | -105.3228
13.8078 | -105.3314
13.8115 | -105.3388
13.8151 | -105.3467
13.8188 | -105.3528
13.8225 | -105.358
13.8262 | -105.3647
13.8298 | -105.3694
13.8335 | -105.37
13.8372 | -105.371
13.8409 | -105.3725
13.8445 | -105.3711
13.8482 | -105.3645
13.8519 | -105.3557
13.8555 | -105.3525
13.8592 | -105.3476
13.8629 | -105.3376
13.8666 | -105.3281
13.8702 | -105.3152
13.8739 | -105.303
13.8776 | -105.2916
13.8813 | -105.2767
13.8849 | -105.26
13.8886 | -105.247
13.8923 | -105.2334
13.896 | -105.2183
13.8996 | -105.2016
13.9033 | -105.1839
13.907 | -105.1658
13.9106 | -105.1475
13.9143 | -105.1323
13.918 | -105.1172
13.9217 | -105.1039
13.9253 | -105.094
13.929 | -105.0846
13.9327 | -105.0724
13.9364 | -105.0632
13.94 | -105.0562
13.9437 | -105.0525
13.9474 | -105.0502
13.9511 | -105.0505
13.9547 | -105.0498
13.9584 | -105.0435
13.9621 | -105.0391
13.9657 | -105.0315
13.9694 | -105.03
13.9731 | -105.0304
13.9768 | -105.0317
13.9804 | -105.0346
13.9841 | -105.0354
13.9878 | -105.0378
13.9915 | -105.0402
13.9951 | -105.0404
13.9988 | -105.0413
14.0025 | -105.0424
14.0062 | -105.0378
14.0098 | -105.0321
14.0135 | -105.0278
14.0172 | -105.0192
14.0208 | -105.0067
14.0245 | -104.9944
14.0282 | -104.9801
14.0319 | -104.9681
14.0355 | -104.9559
14.0392 | -104.9403
14.0429 | -104.9237
14.0466 | -104.9103
14.0502 | -104.9021
14.0539 | -104.894
14.0576 | -104.8825
14.0613 | -104.8732
14.0649 | -104.8693
14.0686 | -104.866
14.0723 | -104.8633
14.0759 | -104.8634
14.0796 | -104.8657
14.0833 | -104.8699
14.087 | -104.8786
14.0906 | -104.887
14.0943 | -104.8987
14.098 | -104.9104
14.1017 | -104.9222
14.1053 | -104.9332
14.109 | -104.9484
14.1127 | -104.9609
14.1164 | -104.9693
14.12 | -104.9772
14.1237 | -104.9832
14.1274 | -104.9884
14.131 | -104.9964
14.1347 | -105.0074
14.1384 | -105.0175
14.1421 | -105.0238
14.1457 | -105.0286
14.1494 | -105.0314
14.1531 | -105.0334
14.1568 | -105.0397
14.1604 | -105.0454
14.1641 | -105.0497
14.1678 | -105.0576
14.1715 | -105.059
14.1751 | -105.0524
14.1788 | -105.0437
14.1825 | -105.0382
14.1861 | -105.0327
14.1898 | -105.0269
14.1935 | -105.0227
14.1972 | -105.0175
14.2008 | -105.0118
14.2045 | -105.0052
14.2082 | -104.9996
14.2119 | -104.9917
14.2155 | -104.9786
14.2192 | -104.9659
14.2229 | -104.9537
14.2265 | -104.9395
14.2302 | -104.9245
14.2339 | -104.9074
14.2376 | -104.8932
14.2412 | -104.8786
14.2449 | -104.8668
14.2486 | -104.8562
14.2523 | -104.8428
14.2559 | -104.8331
14.2596 | -104.8278
14.2633 | -104.8232
14.267 | -104.8187
14.2706 | -104.8175
14.2743 | -104.8201
14.278 | -104.823
14.2816 | -104.8283
14.2853 | -104.8382
14.289 | -104.8466
14.2927 | -104.8569
14.2963 | -104.8683
14.3 | -104.8834
14.3037 | -104.9011
14.3074 | -104.9249
14.311 | -104.9418
14.3147 | -104.9593
14.3184 | -104.9796
14.3221 | -105.0023
14.3257 | -105.0253
14.3294 | -105.0471
14.3331 | -105.067
14.3367 | -105.0859
14.3404 | -105.1051
14.3441 | -105.124
14.3478 | -105.1427
14.3514 | -105.1586
14.3551 | -105.1774
14.3588 | -105.199
14.3625 | -105.214
14.3661 | -105.2284
14.3698 | -105.2417
14.3735 | -105.2499
14.3772 | -105.2575
14.3808 | -105.2682
14.3845 | -105.2775
14.3882 | -105.2817
14.3918 | -105.2853
14.3955 | -105.2862
14.3992 | -105.284
14.4029 | -105.2843
14.4065 | -105.2845
14.4102 | -105.2839
14.4139 | -105.2839
14.4176 | -105.2816
14.4212 | -105.283
14.4249 | -105.2857
14.4286 | -105.2853
14.4323 | -105.2848
14.4359 | -105.2845
14.4396 | -105.2845
14.4433 | -105.2833
14.4469 | -105.2819
14.4506 | -105.2796
14.4543 | -105.2768
14.458 | -105.2775
14.4616 | -105.2746
14.4653 | -105.2681
14.469 | -105.2632
14.4727 | -105.2585
14.4763 | -105.2534
14.48 | -105.2512
14.4837 | -105.247
14.4873 | -105.2417
14.491 | -105.2364
14.4947 | -105.2299
14.4984 | -105.2256
14.502 | -105.2208
14.5057 | -105.2125
14.5094 | -105.2037
14.5131 | -105.1947
14.5167 | -105.1855
14.5204 | -105.1786
14.5241 | -105.1701
14.5278 | -105.1598
14.5314 | -105.1465
14.5351 | -105.1333
14.5388 | -105.1223
14.5424 | -105.1123
14.5461 | -105.1034
14.5498 | -105.0959
14.5535 | -105.09
14.5571 | -105.0856
14.5608 | -105.0825
14.5645 | -105.076
14.5682 | -105.0706
14.5718 | -105.0705
14.5755 | -105.0735
14.5792 | -105.0764
14.5829 | -105.081
14.5865 | -105.0844
14.5902 | -105.0902
14.5939 | -105.096
14.5975 | -105.1025
14.6012 | -105.1113
14.6049 | -105.1203
14.6086 | -105.132
14.6122 | -105.1426
14.6159 | -105.1517
14.6196 | -105.1647
14.6233 | -105.1767
14.6269 | -105.1901
14.6306 | -105.1998
14.6343 | -105.2102
14.638 | -105.2244
14.6416 | -105.2392
14.6453 | -105.2527
14.649 | -105.2628
14.6526 | -105.2703
14.6563 | -105.2796
14.66 | -105.2857
14.6637 | -105.2915
14.6673 | -105.2958
14.671 | -105.3003
14.6747 | -105.3037
14.6784 | -105.3038
14.682 | -105.3064
14.6857 | -105.3108
14.6894 | -105.3108
14.6931 | -105.3096
14.6967 | -105.3066
14.7004 | -105.3005
14.7041 | -105.2989
14.7077 | -105.2944
14.7114 | -105.2893
14.7151 | -105.2885
14.7188 | -105.2842
14.7224 | -105.2788
14.7261 | -105.2743
14.7298 | -105.2684
14.7335 | -105.2621
14.7371 | -105.2566
14.7408 | -105.2549
14.7445 | -105.2528
14.7482 | -105.2457
14.7518 | -105.2359
14.7555 | -105.2266
14.7592 | -105.2165
14.7628 | -105.2053
14.7665 | -105.194
14.7702 | -105.1853
14.7739 | -105.1764
14.7775 | -105.1649
14.7812 | -105.1526
14.7849 | -105.1401
14.7886 | -105.1281
14.7922 | -105.119
14.7959 | -105.1101
14.7996 | -105.0984
14.8033 | -105.0879
14.8069 | -105.0784
14.8106 | -105.0704
14.8143 | -105.0631
14.8179 | -105.0562
14.8216 | -105.05
14.8253 | -105.0423
14.829 | -105.0362
14.8326 | -105.0321
14.8363 | -105.0344
14.84 | -105.0341
14.8437 | -105.035
14.8473 | -105.0341
14.851 | -105.0331
14.8547 | -105.0332
14.8583 | -105.0329
14.862 | -105.0334
14.8657 | -105.0343
14.8694 | -105.0351
14.873 | -105.0395
14.8767 | -105.0439
14.8804 | -105.0479
14.8841 | -105.0496
14.8877 | -105.0529
14.8914 | -105.0574
14.8951 | -105.0591
14.8988 | -105.0582
14.9024 | -105.0578
14.9061 | -105.0569
14.9098 | -105.0592
14.9134 | -105.0656
14.9171 | -105.0701
14.9208 | -105.0733
14.9245 | -105.0783
14.9281 | -105.0826
14.9318 | -105.0882
14.9355 | -105.0936
14.9392 | -105.0973
14.9428 | -105.1033
14.9465 | -105.109
14.9502 | -105.1144
14.9539 | -105.1163
14.9575 | -105.1154
14.9612 | -105.1184
14.9649 | -105.1241
14.9685 | -105.1268
14.9722 | -105.1301
14.9759 | -105.1352
14.9796 | -105.1391
14.9832 | -105.139
14.9869 | -105.1385
14.9906 | -105.1386
14.9943 | -105.1423
14.9979 | -105.1455
15.0016 | -105.1468
15.0053 | -105.1504
15.009 | -105.1572
15.0126 | -105.1673
15.0163 | -105.1757
15.02 | -105.1848
15.0236 | -105.1942
15.0273 | -105.2041
15.031 | -105.2117
15.0347 | -105.2204
15.0383 | -105.2342
15.042 | -105.2468
15.0457 | -105.2612
15.0494 | -105.2748
15.053 | -105.2859
15.0567 | -105.2956
15.0604 | -105.3048
15.0641 | -105.3135
15.0677 | -105.3202
15.0714 | -105.3253
15.0751 | -105.3343
15.0787 | -105.3421
15.0824 | -105.3462
15.0861 | -105.3481
15.0898 | -105.3462
15.0934 | -105.3434
15.0971 | -105.3398
15.1008 | -105.3358
15.1045 | -105.3294
15.1081 | -105.3215
15.1118 | -105.3149
15.1155 | -105.3081
15.1192 | -105.2966
15.1228 | -105.2863
15.1265 | -105.2737
15.1302 | -105.2552
15.1338 | -105.2367
15.1375 | -105.2193
15.1412 | -105.2021
15.1449 | -105.1797
15.1485 | -105.1599
15.1522 | -105.1448
15.1559 | -105.1301
15.1596 | -105.1135
15.1632 | -105.0982
15.1669 | -105.0878
15.1706 | -105.0789
15.1743 | -105.0731
15.1779 | -105.0669
15.1816 | -105.0605
15.1853 | -105.054
15.1889 | -105.0513
15.1926 | -105.0493
15.1963 | -105.0506
15.2 | -105.0533
15.2036 | -105.0583
15.2073 | -105.064
15.211 | -105.0663
15.2147 | -105.0702
15.2183 | -105.0767
15.222 | -105.0825
15.2257 | -105.0885
15.2293 | -105.0932
15.233 | -105.1003
15.2367 | -105.1038
15.2404 | -105.1082
15.244 | -105.1155
15.2477 | -105.1229
15.2514 | -105.1267
15.2551 | -105.1292
15.2587 | -105.1299
15.2624 | -105.1294
15.2661 | -105.1282
15.2698 | -105.1276
15.2734 | -105.1279
15.2771 | -105.1267
15.2808 | -105.1244
15.2844 | -105.1204
15.2881 | -105.1155
15.2918 | -105.1089
15.2955 | -105.1042
15.2991 | -105.0992
15.3028 | -105.0897
15.3065 | -105.0799
15.3102 | -105.0723
15.3138 | -105.066
15.3175 | -105.0623
15.3212 | -105.0576
15.3249 | -105.0518
15.3285 | -105.0425
15.3322 | -105.0354
15.3359 | -105.0292
15.3395 | -105.022
15.3432 | -105.0162
15.3469 | -105.0126
15.3506 | -105.0094
15.3542 | -105.0066
15.3579 | -105.0046
15.3616 | -105.007
15.3653 | -105.0071
15.3689 | -105.0065
15.3726 | -105.0038
15.3763 | -105.0021
15.38 | -105.0023
15.3836 | -105.0022
15.3873 | -104.9985
15.391 | -104.9965
15.3946 | -104.9975
15.3983 | -105.0017
15.402 | -105.0063
15.4057 | -105.0127
15.4093 | -105.0185
15.413 | -105.0274
15.4167 | -105.0339
15.4204 | -105.0431
15.424 | -105.0537
15.4277 | -105.0646
15.4314 | -105.0779
15.4351 | -105.0912
15.4387 | -105.106
15.4424 | -105.1202
15.4461 | -105.1334
15.4497 | -105.1484
15.4534 | -105.1632
15.4571 | -105.1792
15.4608 | -105.1928
15.4644 | -105.2031
15.4681 | -105.2138
15.4718 | -105.2267
15.4755 | -105.2372
15.4791 | -105.2469
15.4828 | -105.2549
15.4865 | -105.2627
15.4902 | -105.2688
15.4938 | -105.2746
15.4975 | -105.2779
15.5012 | -105.2786
15.5048 | -105.2789
15.5085 | -105.2763
15.5122 | -105.2695
15.5159 | -105.2628
15.5195 | -105.2553
15.5232 | -105.2498
15.5269 | -105.2434
15.5306 | -105.2345
15.5342 | -105.2236
15.5379 | -105.2098
15.5416 | -105.1968
15.5452 | -105.1833
15.5489 | -105.1739
15.5526 | -105.163
15.5563 | -105.1531
15.5599 | -105.1446
15.5636 | -105.1334
15.5673 | -105.1221
15.571 | -105.1128
15.5746 | -105.1046
15.5783 | -105.0949
15.582 | -105.0892
15.5857 | -105.0829
15.5893 | -105.0798
15.593 | -105.0792
15.5967 | -105.0788
15.6003 | -105.0782
15.604 | -105.0813
15.6077 | -105.0861
15.6114 | -105.0915
15.615 | -105.0972
15.6187 | -105.1035
15.6224 | -105.1094
15.6261 | -105.1173
15.6297 | -105.1273
15.6334 | -105.1378
15.6371 | -105.1492
15.6408 | -105.1654
15.6444 | -105.1807
15.6481 | -105.1948
15.6518 | -105.2083
15.6554 | -105.2232
15.6591 | -105.2374
15.6628 | -105.2507
15.6665 | -105.2626
15.6701 | -105.2762
15.6738 | -105.2882
15.6775 | -105.3003
15.6812 | -105.3144
15.6848 | -105.3259
15.6885 | -105.3359
15.6922 | -105.3437
15.6959 | -105.3511
15.6995 | -105.3581
15.7032 | -105.3632
15.7069 | -105.3661
15.7105 | -105.367
15.7142 | -105.3696
15.7179 | -105.3692
15.7216 | -105.3674
15.7252 | -105.3634
15.7289 | -105.3595
15.7326 | -105.3547
15.7363 | -105.348
15.7399 | -105.3362
15.7436 | -105.3238
15.7473 | -105.3106
15.751 | -105.2938
15.7546 | -105.2764
15.7583 | -105.2582
15.762 | -105.2395
15.7656 | -105.2193
15.7693 | -105.1988
15.773 | -105.1775
15.7767 | -105.1572
15.7803 | -105.1366
15.784 | -105.1097
15.7877 | -105.0865
15.7914 | -105.064
15.795 | -105.0411
15.7987 | -105.0212
15.8024 | -105.003
15.8061 | -104.985
15.8097 | -104.9707
15.8134 | -104.9548
15.8171 | -104.9399
15.8207 | -104.9249
15.8244 | -104.9115
15.8281 | -104.8974
15.8318 | -104.8877
15.8354 | -104.8817
15.8391 | -104.8761
15.8428 | -104.867
15.8465 | -104.8567
15.8501 | -104.8472
15.8538 | -104.8408
15.8575 | -104.8365
15.8612 | -104.8334
15.8648 | -104.8304
15.8685 | -104.8296
15.8722 | -104.8279
15.8758 | -104.8237
15.8795 | -104.821
15.8832 | -104.8206
15.8869 | -104.8196
15.8905 | -104.8188
15.8942 | -104.8144
15.8979 | -104.8062
15.9016 | -104.8004
15.9052 | -104.7964
15.9089 | -104.7908
15.9126 | -104.7846
15.9162 | -104.781
15.9199 | -104.7767
15.9236 | -104.7657
15.9273 | -104.7556
15.9309 | -104.7458
15.9346 | -104.7374
15.9383 | -104.7265
15.942 | -104.7162
15.9456 | -104.706
15.9493 | -104.6974
15.953 | -104.6896
15.9567 | -104.6774
15.9603 | -104.6632
15.964 | -104.6516
15.9677 | -104.6435
15.9713 | -104.6365
15.975 | -104.6251
15.9787 | -104.6119
15.9824 | -104.6025
15.986 | -104.5948
15.9897 | -104.5894
15.9934 | -104.5821
15.9971 | -104.5768
16.0007 | -104.571
16.0044 | -104.5655
16.0081 | -104.5631
16.0118 | -104.563
16.0154 | -104.5603
16.0191 | -104.5589
16.0228 | -104.5563
16.0264 | -104.5532
16.0301 | -104.5514
16.0338 | -104.5476
16.0375 | -104.5449
16.0411 | -104.5462
16.0448 | -104.5485
16.0485 | -104.5533
16.0522 | -104.5589
16.0558 | -104.563
16.0595 | -104.5667
16.0632 | -104.5701
16.0669 | -104.5758
16.0705 | -104.585
16.0742 | -104.5968
16.0779 | -104.6084
16.0815 | -104.6185
16.0852 | -104.6279
16.0889 | -104.6372
16.0926 | -104.6498
16.0962 | -104.664
16.0999 | -104.6798
16.1036 | -104.6982
16.1073 | -104.7147
16.1109 | -104.7277
16.1146 | -104.7412
16.1183 | -104.7568
16.122 | -104.7741
16.1256 | -104.7883
16.1293 | -104.8035
16.133 | -104.8185
16.1366 | -104.8284
16.1403 | -104.8375
16.144 | -104.8464
16.1477 | -104.8508
16.1513 | -104.8546
16.155 | -104.8627
16.1587 | -104.8691
16.1624 | -104.8753
16.166 | -104.8811
16.1697 | -104.8873
16.1734 | -104.8901
16.1771 | -104.8905
16.1807 | -104.8919
16.1844 | -104.8919
16.1881 | -104.8943
16.1917 | -104.8993
16.1954 | -104.9021
16.1991 | -104.9011
16.2028 | -104.9008
16.2064 | -104.9047
16.2101 | -104.9109
16.2138 | -104.911
16.2175 | -104.9102
16.2211 | -104.9114
16.2248 | -104.9164
16.2285 | -104.9208
16.2322 | -104.924
16.2358 | -104.9255
16.2395 | -104.9268
16.2432 | -104.9325
16.2468 | -104.9382
16.2505 | -104.9452
16.2542 | -104.9521
16.2579 | -104.958
16.2615 | -104.9623
16.2652 | -104.9679
16.2689 | -104.9734
16.2726 | -104.9751
16.2762 | -104.976
16.2799 | -104.9758
16.2836 | -104.9762
16.2872 | -104.9762
16.2909 | -104.9764
16.2946 | -104.9792
16.2983 | -104.9789
16.3019 | -104.9795
16.3056 | -104.9791
16.3093 | -104.98
16.313 | -104.982
16.3166 | -104.9829
16.3203 | -104.9856
16.324 | -104.9882
16.3277 | -104.9895
16.3313 | -104.992
16.335 | -104.993
16.3387 | -104.9929
16.3423 | -104.9993
16.346 | -105.0041
16.3497 | -105.006
16.3534 | -105.0074
16.357 | -105.0107
16.3607 | -105.0132
16.3644 | -105.0131
16.3681 | -105.0134
16.3717 | -105.0157
16.3754 | -105.0195
16.3791 | -105.024
16.3828 | -105.0278
16.3864 | -105.0295
16.3901 | -105.0284
16.3938 | -105.0283
16.3974 | -105.0276
16.4011 | -105.028
16.4048 | -105.0294
16.4085 | -105.0341
16.4121 | -105.0357
16.4158 | -105.0403
16.4195 | -105.0434
16.4232 | -105.047
16.4268 | -105.0583
16.4305 | -105.0654
16.4342 | -105.0748
16.4379 | -105.0818
16.4415 | -105.0904
16.4452 | -105.0985
16.4489 | -105.1095
16.4525 | -105.1234
16.4562 | -105.1356
16.4599 | -105.1483
16.4636 | -105.162
16.4672 | -105.1765
16.4709 | -105.1932
16.4746 | -105.21
16.4783 | -105.2233
16.4819 | -105.2351
16.4856 | -105.245
16.4893 | -105.2559
16.493 | -105.2666
16.4966 | -105.2775
16.5003 | -105.2866
16.504 | -105.2912
16.5076 | -105.2949
16.5113 | -105.2987
16.515 | -105.3018
16.5187 | -105.3014
16.5223 | -105.2998
16.526 | -105.2905
16.5297 | -105.2794
16.5334 | -105.2728
16.537 | -105.2641
16.5407 | -105.25
16.5444 | -105.233
16.5481 | -105.2147
16.5517 | -105.198
16.5554 | -105.1799
16.5591 | -105.1621
16.5627 | -105.1478
16.5664 | -105.1356
16.5701 | -105.125
16.5738 | -105.1144
16.5774 | -105.1001
16.5811 | -105.0892
16.5848 | -105.0814
16.5885 | -105.075
16.5921 | -105.0659
16.5958 | -105.0613
16.5995 | -105.0602
16.6031 | -105.0595
16.6068 | -105.0578
16.6105 | -105.0568
16.6142 | -105.0579
16.6178 | -105.0602
16.6215 | -105.0637
16.6252 | -105.0711
16.6289 | -105.077
16.6325 | -105.0861
16.6362 | -105.0956
16.6399 | -105.1071
16.6436 | -105.1203
16.6472 | -105.1343
16.6509 | -105.1483
16.6546 | -105.1658
16.6582 | -105.1834
16.6619 | -105.2008
16.6656 | -105.2187
16.6693 | -105.2363
16.6729 | -105.2538
16.6766 | -105.2722
16.6803 | -105.2907
16.684 | -105.309
16.6876 | -105.3229
16.6913 | -105.335
16.695 | -105.346
16.6987 | -105.3585
16.7023 | -105.3685
16.706 | -105.379
16.7097 | -105.3886
16.7133 | -105.3945
16.717 | -105.3988
16.7207 | -105.4045
16.7244 | -105.4089
16.728 | -105.4111
16.7317 | -105.4113
16.7354 | -105.4099
16.7391 | -105.4089
16.7427 | -105.4092
16.7464 | -105.4083
16.7501 | -105.4038
16.7538 | -105.396
16.7574 | -105.3874
16.7611 | -105.3814
16.7648 | -105.3779
16.7684 | -105.3711
16.7721 | -105.3632
16.7758 | -105.3539
16.7795 | -105.3465
16.7831 | -105.3408
16.7868 | -105.3361
16.7905 | -105.332
16.7942 | -105.3272
16.7978 | -105.3208
16.8015 | -105.3136
16.8052 | -105.3066
16.8089 | -105.3023
16.8125 | -105.2996
16.8162 | -105.299
16.8199 | -105.2985
16.8235 | -105.2982
16.8272 | -105.296
16.8309 | -105.297
16.8346 | -105.2975
16.8382 | -105.2973
16.8419 | -105.2992
16.8456 | -105.3034
16.8493 | -105.3062
16.8529 | -105.3081
16.8566 | -105.3134
16.8603 | -105.3193
16.864 | -105.3222
16.8676 | -105.3259
16.8713 | -105.3276
16.875 | -105.3301
16.8786 | -105.3296
16.8823 | -105.3268
16.886 | -105.3243
16.8897 | -105.32
16.8933 | -105.3155
16.897 | -105.3092
16.9007 | -105.3021
16.9044 | -105.2911
16.908 | -105.2792
16.9117 | -105.268
16.9154 | -105.255
16.919 | -105.2394
16.9227 | -105.2251
16.9264 | -105.2109
16.9301 | -105.1956
16.9337 | -105.1767
16.9374 | -105.1534
16.9411 | -105.1315
16.9448 | -105.1093
16.9484 | -105.0863
16.9521 | -105.0653
16.9558 | -105.048
16.9595 | -105.0294
16.9631 | -105.0109
16.9668 | -104.9947
16.9705 | -104.9781
16.9741 | -104.9633
16.9778 | -104.9491
16.9815 | -104.9358
16.9852 | -104.9228
16.9888 | -104.9122
16.9925 | -104.9053
16.9962 | -104.902
16.9999 | -104.9021
17.0035 | -104.9012
17.0072 | -104.9012
17.0109 | -104.9042
17.0146 | -104.9093
17.0182 | -104.9167
17.0219 | -104.926
17.0256 | -104.9375
17.0292 | -104.9465
17.0329 | -104.9572
17.0366 | -104.9713
17.0403 | -104.9868
17.0439 | -105.0029
17.0476 | -105.0209
17.0513 | -105.0388
17.055 | -105.0578
17.0586 | -105.0777
17.0623 | -105.0999
17.066 | -105.1199
17.0697 | -105.1416
17.0733 | -105.1623
17.077 | -105.1823
17.0807 | -105.2024
17.0843 | -105.2201
17.088 | -105.2389
17.0917 | -105.2589
17.0954 | -105.2784
17.099 | -105.2934
17.1027 | -105.3081
17.1064 | -105.3238
17.1101 | -105.3382
17.1137 | -105.3518
17.1174 | -105.3631
17.1211 | -105.3758
17.1248 | -105.3891
17.1284 | -105.3982
17.1321 | -105.4069
17.1358 | -105.4146
17.1394 | -105.4224
17.1431 | -105.4272
17.1468 | -105.4308
17.1505 | -105.434
17.1541 | -105.4347
17.1578 | -105.4348
17.1615 | -105.4338
17.1652 | -105.4311
17.1688 | -105.4277
17.1725 | -105.4209
17.1762 | -105.4143
17.1799 | -105.4083
17.1835 | -105.3988
17.1872 | -105.3885
17.1909 | -105.379
17.1945 | -105.3691
17.1982 | -105.3568
17.2019 | -105.3422
17.2056 | -105.3273
17.2092 | -105.3115
17.2129 | -105.2935
17.2166 | -105.2788
17.2203 | -105.2656
17.2239 | -105.2515
17.2276 | -105.2372
17.2313 | -105.2244
17.235 | -105.2115
17.2386 | -105.1975
17.2423 | -105.1844
17.246 | -105.1707
17.2496 | -105.1583
17.2533 | -105.1466
17.257 | -105.1315
17.2607 | -105.1181
17.2643 | -105.1092
17.268 | -105.1036
17.2717 | -105.0977
17.2754 | -105.0947
17.279 | -105.0898
17.2827 | -105.0845
17.2864 | -105.0813
17.2901 | -105.0796
17.2937 | -105.0797
17.2974 | -105.0828
17.3011 | -105.0863
17.3047 | -105.0887
17.3084 | -105.0891
17.3121 | -105.0915
17.3158 | -105.0963
17.3194 | -105.1017
17.3231 | -105.1082
17.3268 | -105.1135
17.3305 | -105.1177
17.3341 | -105.1196
17.3378 | -105.122
17.3415 | -105.1229
17.3451 | -105.1262
17.3488 | -105.1286
17.3525 | -105.1313
17.3562 | -105.1343
17.3598 | -105.1371
17.3635 | -105.1412
17.3672 | -105.1428
17.3709 | -105.1431
17.3745 | -105.1467
17.3782 | -105.1494
17.3819 | -105.1486
17.3856 | -105.1462
17.3892 | -105.1472
17.3929 | -105.1476
17.3966 | -105.148
17.4002 | -105.1491
17.4039 | -105.1483
17.4076 | -105.1478
17.4113 | -105.1427
17.4149 | -105.1402
17.4186 | -105.1398
17.4223 | -105.1384
17.426 | -105.1359
17.4296 | -105.134
17.4333 | -105.1295
17.437 | -105.12
17.4407 | -105.1101
17.4443 | -105.1046
17.448 | -105.0985
17.4517 | -105.0924
17.4553 | -105.0836
17.459 | -105.0752
17.4627 | -105.0638
17.4664 | -105.0498
17.47 | -105.035
17.4737 | -105.0209
17.4774 | -105.0068
17.4811 | -104.993
17.4847 | -104.9774
17.4884 | -104.9643
17.4921 | -104.9512
17.4958 | -104.9384
17.4994 | -104.9253
17.5031 | -104.912
17.5068 | -104.9016
17.5104 | -104.8891
17.5141 | -104.8759
17.5178 | -104.8631
17.5215 | -104.8519
17.5251 | -104.8434
17.5288 | -104.8367
17.5325 | -104.8314
17.5362 | -104.8263
17.5398 | -104.8214
17.5435 | -104.8167
17.5472 | -104.8112
17.5509 | -104.8071
17.5545 | -104.8056
17.5582 | -104.8069
17.5619 | -104.8113
17.5655 | -104.8142
17.5692 | -104.8179
17.5729 | -104.8212
17.5766 | -104.8262
17.5802 | -104.8333
17.5839 | -104.8392
17.5876 | -104.8446
17.5913 | -104.854
17.5949 | -104.862
17.5986 | -104.8684
17.6023 | -104.8731
17.606 | -104.878
17.6096 | -104.884
17.6133 | -104.8887
17.617 | -104.893
17.6206 | -104.8979
17.6243 | -104.9007
17.628 | -104.9035
17.6317 | -104.9085
17.6353 | -104.9139
17.639 | -104.9172
17.6427 | -104.9199
17.6464 | -104.9213
17.65 | -104.9209
17.6537 | -104.9232
17.6574 | -104.9206
17.661 | -104.9204
17.6647 | -104.9164
17.6684 | -104.9113
17.6721 | -104.905
17.6757 | -104.8971
17.6794 | -104.8897
17.6831 | -104.8827
17.6868 | -104.8762
17.6904 | -104.8686
17.6941 | -104.8577
17.6978 | -104.8467
17.7015 | -104.8355
17.7051 | -104.8245
17.7088 | -104.8138
17.7125 | -104.8039
17.7161 | -104.7938
17.7198 | -104.7851
17.7235 | -104.7762
17.7272 | -104.7679
17.7308 | -104.7583
17.7345 | -104.7503
17.7382 | -104.7427
17.7419 | -104.7362
17.7455 | -104.731
17.7492 | -104.7269
17.7529 | -104.7228
17.7566 | -104.7199
17.7602 | -104.7176
17.7639 | -104.7171
17.7676 | -104.7164
17.7712 | -104.7143
17.7749 | -104.7135
17.7786 | -104.7138
17.7823 | -104.7152
17.7859 | -104.7179
17.7896 | -104.7222
17.7933 | -104.7267
17.797 | -104.7305
17.8006 | -104.7368
17.8043 | -104.7434
17.808 | -104.7504
17.8117 | -104.7568
17.8153 | -104.7624
17.819 | -104.7684
17.8227 | -104.7756
17.8263 | -104.7844
17.83 | -104.7931
17.8337 | -104.7996
17.8374 | -104.8086
17.841 | -104.8193
17.8447 | -104.8272
17.8484 | -104.8358
17.8521 | -104.8473
17.8557 | -104.8586
17.8594 | -104.8683
17.8631 | -104.8771
17.8668 | -104.8836
17.8704 | -104.8927
17.8741 | -104.8998
17.8778 | -104.9077
17.8814 | -104.9157
17.8851 | -104.92
17.8888 | -104.9251
17.8925 | -104.9274
17.8961 | -104.9293
17.8998 | -104.9306
17.9035 | -104.9312
17.9072 | -104.9316
17.9108 | -104.9343
17.9145 | -104.9338
17.9182 | -104.9304
17.9219 | -104.9273
17.9255 | -104.9268
17.9292 | -104.9261
17.9329 | -104.9214
17.9365 | -104.9131
17.9402 | -104.9056
17.9439 | -104.8962
17.9476 | -104.8878
17.9512 | -104.8809
17.9549 | -104.8734
17.9586 | -104.8643
17.9623 | -104.8526
17.9659 | -104.8391
17.9696 | -104.8261
17.9733 | -104.8141
17.9769 | -104.8012
17.9806 | -104.7866
17.9843 | -104.7742
17.988 | -104.7611
17.9916 | -104.7464
17.9953 | -104.7324
17.999 | -104.7183
18.0027 | -104.7018
18.0063 | -104.6853
18.01 | -104.6725
18.0137 | -104.657
18.0174 | -104.6399
18.021 | -104.6236
18.0247 | -104.6075
18.0284 | -104.594
18.032 | -104.5785
18.0357 | -104.564
18.0394 | -104.5472
18.0431 | -104.5302
18.0467 | -104.5162
18.0504 | -104.5046
18.0541 | -104.495
18.0578 | -104.4841
18.0614 | -104.4745
18.0651 | -104.4637
18.0688 | -104.4535
18.0725 | -104.4456
18.0761 | -104.4407
18.0798 | -104.4354
18.0835 | -104.4291
18.0871 | -104.4241
18.0908 | -104.4208
18.0945 | -104.4143
18.0982 | -104.4077
18.1018 | -104.4028
18.1055 | -104.3983
18.1092 | -104.3944
18.1129 | -104.3932
18.1165 | -104.3924
18.1202 | -104.3897
18.1239 | -104.3892
18.1276 | -104.3868
18.1312 | -104.3859
18.1349 | -104.3851
18.1386 | -104.384
18.1422 | -104.3841
18.1459 | -104.3843
18.1496 | -104.3867
18.1533 | -104.3897
18.1569 | -104.39
18.1606 | -104.3889
18.1643 | -104.3864
18.168 | -104.384
18.1716 | -104.3855
18.1753 | -104.389
18.179 | -104.3905
18.1827 | -104.3917
18.1863 | -104.3932
18.19 | -104.3956
18.1937 | -104.4017
18.1973 | -104.4086
18.201 | -104.4126
18.2047 | -104.4158
18.2084 | -104.4203
18.212 | -104.4267
18.2157 | -104.4336
18.2194 | -104.4411
18.2231 | -104.4488
18.2267 | -104.4555
18.2304 | -104.4599
18.2341 | -104.4661
18.2378 | -104.474
18.2414 | -104.4821
18.2451 | -104.4877
18.2488 | -104.4941
18.2524 | -104.5017
18.2561 | -104.5085
18.2598 | -104.5171
18.2635 | -104.5256
18.2671 | -104.53
18.2708 | -104.5334
18.2745 | -104.5369
18.2782 | -104.5404
18.2818 | -104.5427
18.2855 | -104.5452
18.2892 | -104.5478
18.2929 | -104.5487
18.2965 | -104.5449
18.3002 | -104.5438
18.3039 | -104.5426
18.3075 | -104.5394
18.3112 | -104.5327
18.3149 | -104.5247
18.3186 | -104.5182
18.3222 | -104.51
18.3259 | -104.5017
18.3296 | -104.4922
18.3333 | -104.4817
18.3369 | -104.4699
18.3406 | -104.4571
18.3443 | -104.4456
18.348 | -104.4336
18.3516 | -104.4209
18.3553 | -104.4111
18.359 | -104.4015
18.3626 | -104.3904
18.3663 | -104.3796
18.37 | -104.3682
18.3737 | -104.3546
18.3773 | -104.3438
18.381 | -104.3355
18.3847 | -104.3252
18.3884 | -104.3163
18.392 | -104.3061
18.3957 | -104.2973
18.3994 | -104.2888
18.403 | -104.2803
18.4067 | -104.2729
18.4104 | -104.2654
18.4141 | -104.2569
18.4177 | -104.2514
18.4214 | -104.2474
18.4251 | -104.2438
18.4288 | -104.2418
18.4324 | -104.2371
18.4361 | -104.2317
18.4398 | -104.229
18.4435 | -104.2257
18.4471 | -104.2224
18.4508 | -104.2213
18.4545 | -104.2218
18.4581 | -104.2237
18.4618 | -104.229
18.4655 | -104.2337
18.4692 | -104.2368
18.4728 | -104.24
18.4765 | -104.247
18.4802 | -104.2543
18.4839 | -104.2618
18.4875 | -104.2696
18.4912 | -104.277
18.4949 | -104.2828
18.4986 | -104.2897
18.5022 | -104.2965
18.5059 | -104.305
18.5096 | -104.314
18.5132 | -104.3212
18.5169 | -104.3291
18.5206 | -104.3369
18.5243 | -104.3459
18.5279 | -104.3555
18.5316 | -104.3661
18.5353 | -104.374
18.539 | -104.3804
18.5426 | -104.387
18.5463 | -104.3908
18.55 | -104.3943
18.5537 | -104.397
18.5573 | -104.4026
18.561 | -104.4117
18.5647 | -104.4187
18.5683 | -104.422
18.572 | -104.4244
18.5757 | -104.4289
18.5794 | -104.4334
18.583 | -104.4377
18.5867 | -104.442
18.5904 | -104.4467
18.5941 | -104.4526
18.5977 | -104.455
18.6014 | -104.4587
18.6051 | -104.4637
18.6088 | -104.4702
18.6124 | -104.4757
18.6161 | -104.4803
18.6198 | -104.4822
18.6234 | -104.4831
18.6271 | -104.4851
18.6308 | -104.4879
18.6345 | -104.4903
18.6381 | -104.4897
18.6418 | -104.4887
18.6455 | -104.4881
18.6492 | -104.4881
18.6528 | -104.4869
18.6565 | -104.4871
18.6602 | -104.4865
18.6639 | -104.485
18.6675 | -104.4817
18.6712 | -104.4761
18.6749 | -104.4743
18.6785 | -104.4708
18.6822 | -104.4676
18.6859 | -104.4614
18.6896 | -104.4535
18.6932 | -104.444
18.6969 | -104.4357
18.7006 | -104.4275
18.7043 | -104.4169
18.7079 | -104.4078
18.7116 | -104.3967
18.7153 | -104.383
18.7189 | -104.3672
18.7226 | -104.3489
18.7263 | -104.3299
18.73 | -104.3103
18.7336 | -104.2927
18.7373 | -104.2748
18.741 | -104.2507
18.7447 | -104.2257
18.7483 | -104.2043
18.752 | -104.1831
18.7557 | -104.1606
18.7594 | -104.1396
18.763 | -104.1147
18.7667 | -104.0904
18.7704 | -104.0649
18.774 | -104.0369
18.7777 | -104.0108
18.7814 | -103.9853
18.7851 | -103.9594
18.7887 | -103.9317
18.7924 | -103.9071
18.7961 | -103.8834
18.7998 | -103.8615
18.8034 | -103.8402
18.8071 | -103.8193
18.8108 | -103.7983
18.8145 | -103.7786
18.8181 | -103.7604
18.8218 | -103.7444
18.8255 | -103.7285
18.8291 | -103.7125
18.8328 | -103.6972
18.8365 | -103.6843
18.8402 | -103.6725
18.8438 | -103.6616
18.8475 | -103.6533
18.8512 | -103.6455
18.8549 | -103.6368
18.8585 | -103.6286
18.8622 | -103.6219
18.8659 | -103.6181
18.8696 | -103.6123
18.8732 | -103.6078
18.8769 | -103.5997
18.8806 | -103.5949
18.8842 | -103.5913
18.8879 | -103.589
18.8916 | -103.5839
18.8953 | -103.581
18.8989 | -103.5792
18.9026 | -103.5792
18.9063 | -103.5821
18.91 | -103.5829
18.9136 | -103.5814
18.9173 | -103.58
18.921 | -103.5775
18.9247 | -103.5741
18.9283 | -103.571
18.932 | -103.5685
18.9357 | -103.5682
18.9393 | -103.5637
18.943 | -103.5554
18.9467 | -103.549
18.9504 | -103.5405
18.954 | -103.5338
18.9577 | -103.5252
18.9614 | -103.5146
18.9651 | -103.5035
18.9687 | -103.4943
18.9724 | -103.4836
18.9761 | -103.4699
18.9798 | -103.4549
18.9834 | -103.44
18.9871 | -103.425
18.9908 | -103.4094
18.9944 | -103.3936
18.9981 | -103.3774
19.0018 | -103.3595
19.0055 | -103.341
19.0091 | -103.3224
19.0128 | -103.3036
19.0165 | -103.2816
19.0202 | -103.2603
19.0238 | -103.2394
19.0275 | -103.2166
19.0312 | -103.1981
19.0348 | -103.1793
19.0385 | -103.1612
19.0422 | -103.1435
19.0459 | -103.1255
19.0495 | -103.1092
19.0532 | -103.0942
19.0569 | -103.0788
19.0606 | -103.063
19.0642 | -103.0497
19.0679 | -103.0376
19.0716 | -103.0255
19.0753 | -103.0215
19.0789 | -103.0113
19.0826 | -103.0027
19.0863 | -102.9962
19.0899 | -102.9908
19.0936 | -102.9838
19.0973 | -102.9763
19.101 | -102.97
19.1046 | -102.9626
19.1083 | -102.956
19.112 | -102.9507
19.1157 | -102.9459
19.1193 | -102.9408
19.123 | -102.9371
19.1267 | -102.9311
19.1304 | -102.9274
19.134 | -102.9215
19.1377 | -102.9139
19.1414 | -102.9048
19.145 | -102.895
19.1487 | -102.8898
19.1524 | -102.882
19.1561 | -102.8738
19.1597 | -102.8634
19.1634 | -102.8544
19.1671 | -102.8427
19.1708 | -102.8299
19.1744 | -102.8162
19.1781 | -102.8029
19.1818 | -102.7882
19.1855 | -102.7726
19.1891 | -102.7592
19.1928 | -102.747
19.1965 | -102.7338
19.2001 | -102.719
19.2038 | -102.7044
19.2075 | -102.6898
19.2112 | -102.6792
19.2148 | -102.6668
19.2185 | -102.6548
19.2222 | -102.6431
19.2259 | -102.632
19.2295 | -102.6233
19.2332 | -102.6165
19.2369 | -102.6087
19.2406 | -102.6005
19.2442 | -102.5953
19.2479 | -102.5937
19.2516 | -102.5954
19.2552 | -102.5964
19.2589 | -102.5977
19.2626 | -102.6004
19.2663 | -102.6051
19.2699 | -102.6127
19.2736 | -102.6211
19.2773 | -102.6306
19.281 | -102.6429
19.2846 | -102.6553
19.2883 | -102.6712
19.292 | -102.6868
19.2957 | -102.7039
19.2993 | -102.722
19.303 | -102.7379
19.3067 | -102.7545
19.3103 | -102.7729
19.314 | -102.7928
19.3177 | -102.8113
19.3214 | -102.827
19.325 | -102.844
19.3287 | -102.863
19.3324 | -102.8795
19.3361 | -102.8927
19.3397 | -102.9043
19.3434 | -102.9165
19.3471 | -102.9284
19.3508 | -102.9386
19.3544 | -102.9465
19.3581 | -102.9551
19.3618 | -102.9622
19.3654 | -102.9659
19.3691 | -102.9686
19.3728 | -102.9702
19.3765 | -102.9695
19.3801 | -102.97
19.3838 | -102.9677
19.3875 | -102.9608
19.3912 | -102.9502
19.3948 | -102.9391
19.3985 | -102.9268
19.4022 | -102.9186
19.4059 | -102.9084
19.4095 | -102.8954
19.4132 | -102.8804
19.4169 | -102.8679
19.4205 | -102.8533
19.4242 | -102.8365
19.4279 | -102.8204
19.4316 | -102.8035
19.4352 | -102.7881
19.4389 | -102.7701
19.4426 | -102.75
19.4463 | -102.7316
19.4499 | -102.7125
19.4536 | -102.6961
19.4573 | -102.6819
19.4609 | -102.668
19.4646 | -102.6512
19.4683 | -102.6345
19.472 | -102.6185
19.4756 | -102.6044
19.4793 | -102.5893
19.483 | -102.574
19.4867 | -102.5589
19.4903 | -102.5442
19.494 | -102.5296
19.4977 | -102.5157
19.5014 | -102.5005
19.505 | -102.4863
19.5087 | -102.4735
19.5124 | -102.4642
19.516 | -102.4547
19.5197 | -102.4446
19.5234 | -102.4341
19.5271 | -102.4241
19.5307 | -102.417
19.5344 | -102.4102
19.5381 | -102.4057
19.5418 | -102.4012
19.5454 | -102.3954
19.5491 | -102.3916
19.5528 | -102.3896
19.5565 | -102.3882
19.5601 | -102.3883
19.5638 | -102.3875
19.5675 | -102.3887
19.5711 | -102.3896
19.5748 | -102.3921
19.5785 | -102.3976
19.5822 | -102.402
19.5858 | -102.4067
19.5895 | -102.4119
19.5932 | -102.4196
19.5969 | -102.4285
19.6005 | -102.4364
19.6042 | -102.4498
19.6079 | -102.4576
19.6116 | -102.467
19.6152 | -102.4783
19.6189 | -102.4911
19.6226 | -102.5018
19.6262 | -102.5128
19.6299 | -102.5229
19.6336 | -102.5362
19.6373 | -102.5492
19.6409 | -102.5624
19.6446 | -102.5771
19.6483 | -102.5907
19.652 | -102.6066
19.6556 | -102.624
19.6593 | -102.6403
19.663 | -102.6579
19.6667 | -102.6762
19.6703 | -102.6946
19.674 | -102.712
19.6777 | -102.7272
19.6813 | -102.7416
19.685 | -102.7558
19.6887 | -102.7713
19.6924 | -102.788
19.696 | -102.8028
19.6997 | -102.8168
19.7034 | -102.8284
19.7071 | -102.8388
19.7107 | -102.849
19.7144 | -102.8598
19.7181 | -102.8729
19.7218 | -102.8849
19.7254 | -102.892
19.7291 | -102.8987
19.7328 | -102.9034
19.7364 | -102.9076
19.7401 | -102.9123
19.7438 | -102.9145
19.7475 | -102.9177
19.7511 | -102.9204
19.7548 | -102.9203
19.7585 | -102.9197
19.7622 | -102.9194
19.7658 | -102.9188
19.7695 | -102.918
19.7732 | -102.9171
19.7768 | -102.9162
19.7805 | -102.9157
19.7842 | -102.9124
19.7879 | -102.9077
19.7915 | -102.8994
19.7952 | -102.8926
19.7989 | -102.8872
19.8026 | -102.8799
19.8062 | -102.872
19.8099 | -102.8629
19.8136 | -102.8526
19.8173 | -102.8418
19.8209 | -102.8315
19.8246 | -102.821
19.8283 | -102.8098
19.8319 | -102.7973
19.8356 | -102.7835
19.8393 | -102.7703
19.843 | -102.757
19.8466 | -102.7435
19.8503 | -102.7302
19.854 | -102.7155
19.8577 | -102.701
19.8613 | -102.6869
19.865 | -102.6722
19.8687 | -102.658
19.8724 | -102.6429
19.876 | -102.6289
19.8797 | -102.6153
19.8834 | -102.6007
19.887 | -102.5873
19.8907 | -102.5772
19.8944 | -102.5678
19.8981 | -102.5574
19.9017 | -102.5467
19.9054 | -102.5388
19.9091 | -102.5317
19.9128 | -102.5238
19.9164 | -102.5141
19.9201 | -102.5067
19.9238 | -102.4996
19.9275 | -102.4929
19.9311 | -102.4893
19.9348 | -102.4864
19.9385 | -102.4839
19.9421 | -102.4828
19.9458 | -102.4806
19.9495 | -102.4809
19.9532 | -102.4818
19.9568 | -102.4819
19.9605 | -102.4848
19.9642 | -102.4881
19.9679 | -102.491
19.9715 | -102.4949
19.9752 | -102.5001
19.9789 | -102.5049
19.9826 | -102.511
19.9862 | -102.518
19.9899 | -102.5248
19.9936 | -102.5305
19.9972 | -102.5359
20.0009 | -102.541
20.0046 | -102.5493
20.0083 | -102.558
20.0119 | -102.568
20.0156 | -102.5766
20.0193 | -102.5861
20.023 | -102.5965
20.0266 | -102.61
20.0303 | -102.6219
20.034 | -102.6333
20.0377 | -102.6475
20.0413 | -102.6617
20.045 | -102.6747
20.0487 | -102.6877
20.0523 | -102.7019
20.056 | -102.7149
20.0597 | -102.7281
20.0634 | -102.7433
20.067 | -102.7606
20.0707 | -102.7778
20.0744 | -102.7953
20.0781 | -102.8126
20.0817 | -102.8281
20.0854 | -102.8456
20.0891 | -102.8635
20.0927 | -102.8786
20.0964 | -102.8934
20.1001 | -102.9093
20.1038 | -102.9237
20.1074 | -102.9375
20.1111 | -102.948
20.1148 | -102.9566
20.1185 | -102.968
20.1221 | -102.9768
20.1258 | -102.9812
20.1295 | -102.9848
20.1332 | -102.9892
20.1368 | -102.9834
20.1405 | -102.9842
20.1442 | -102.9834
20.1478 | -102.9827
20.1515 | -102.9776
20.1552 | -102.9694
20.1589 | -102.9622
20.1625 | -102.9548
20.1662 | -102.9452
20.1699 | -102.9359
20.1736 | -102.9248
20.1772 | -102.9087
20.1809 | -102.8891
20.1846 | -102.8694
20.1883 | -102.8488
20.1919 | -102.829
20.1956 | -102.8071
20.1993 | -102.7877
20.2029 | -102.7675
20.2066 | -102.7427
20.2103 | -102.718
20.214 | -102.6938
20.2176 | -102.6692
20.2213 | -102.6413
20.225 | -102.6127
20.2287 | -102.5831
20.2323 | -102.5529
20.236 | -102.5235
20.2397 | -102.4968
20.2434 | -102.4697
20.247 | -102.4407
20.2507 | -102.4136
20.2544 | -102.3873
20.258 | -102.3623
20.2617 | -102.3362
20.2654 | -102.3113
20.2691 | -102.2866
20.2727 | -102.263
20.2764 | -102.2398
20.2801 | -102.2204
20.2838 | -102.2003
20.2874 | -102.181
20.2911 | -102.1611
20.2948 | -102.1438
20.2985 | -102.1242
20.3021 | -102.1049
20.3058 | -102.087
20.3095 | -102.0701
20.3131 | -102.0536
20.3168 | -102.0387
20.3205 | -102.0255
20.3242 | -102.0109
20.3278 | -101.9965
20.3315 | -101.9835
20.3352 | -101.971
20.3389 | -101.9588
20.3425 | -101.9471
20.3462 | -101.9362
20.3499 | -101.9226
20.3536 | -101.9083
20.3572 | -101.895
20.3609 | -101.8831
20.3646 | -101.871
20.3682 | -101.8587
20.3719 | -101.8469
20.3756 | -101.8355
20.3793 | -101.8267
20.3829 | -101.8193
20.3866 | -101.8085
20.3903 | -101.7991
20.394 | -101.7902
20.3976 | -101.7828
20.4013 | -101.7759
20.405 | -101.7673
20.4087 | -101.7603
20.4123 | -101.7566
20.416 | -101.7525
20.4197 | -101.7463
20.4233 | -101.7374
20.427 | -101.7314
20.4307 | -101.7264
20.4344 | -101.7212
20.438 | -101.7198
20.4417 | -101.7215
20.4454 | -101.7229
20.4491 | -101.72
20.4527 | -101.7153
20.4564 | -101.7136
20.4601 | -101.7114
20.4637 | -101.7101
20.4674 | -101.7104
20.4711 | -101.7117
20.4748 | -101.7119
20.4784 | -101.7115
20.4821 | -101.7096
20.4858 | -101.7095
20.4895 | -101.708
20.4931 | -101.7069
20.4968 | -101.7066
20.5005 | -101.7056
20.5042 | -101.7028
20.5078 | -101.6984
20.5115 | -101.6953
20.5152 | -101.6918
20.5188 | -101.6892
20.5225 | -101.685
20.5262 | -101.6782
20.5299 | -101.6695
20.5335 | -101.6623
20.5372 | -101.6538
20.5409 | -101.6427
20.5446 | -101.6313
20.5482 | -101.6185
20.5519 | -101.6035
20.5556 | -101.5868
20.5593 | -101.5686
20.5629 | -101.5488
20.5666 | -101.5311
20.5703 | -101.5131
20.5739 | -101.492
20.5776 | -101.4689
20.5813 | -101.448
20.585 | -101.4245
20.5886 | -101.4026
20.5923 | -101.3818
20.596 | -101.3583
20.5997 | -101.3323
20.6033 | -101.3039
20.607 | -101.276
20.6107 | -101.2504
20.6144 | -101.2243
20.618 | -101.1984
20.6217 | -101.1734
20.6254 | -101.1426
20.629 | -101.1085
20.6327 | -101.0781
20.6364 | -101.046
20.6401 | -101.0143
20.6437 | -100.9852
20.6474 | -100.9536
20.6511 | -100.9202
20.6548 | -100.8877
20.6584 | -100.8568
20.6621 | -100.8255
20.6658 | -100.793
20.6695 | -100.7619
20.6731 | -100.7274
20.6768 | -100.6945
20.6805 | -100.6636
20.6841 | -100.6334
20.6878 | -100.6038
20.6915 | -100.576
20.6952 | -100.5491
20.6988 | -100.5211
20.7025 | -100.4916
20.7062 | -100.4639
20.7099 | -100.4395
20.7135 | -100.4167
20.7172 | -100.3956
20.7209 | -100.3724
20.7246 | -100.3504
20.7282 | -100.327
20.7319 | -100.3032
20.7356 | -100.281
20.7392 | -100.2592
20.7429 | -100.2396
20.7466 | -100.2212
20.7503 | -100.2032
20.7539 | -100.187
20.7576 | -100.1711
20.7613 | -100.1553
20.765 | -100.1399
20.7686 | -100.1244
20.7723 | -100.1129
20.776 | -100.1019
20.7797 | -100.092
20.7833 | -100.0829
20.787 | -100.0749
20.7907 | -100.0683
20.7943 | -100.0631
20.798 | -100.0575
20.8017 | -100.0528
20.8054 | -100.051
20.809 | -100.0479
20.8127 | -100.045
20.8164 | -100.0456
20.8201 | -100.0479
20.8237 | -100.0511
20.8274 | -100.0535
20.8311 | -100.0565
20.8347 | -100.0596
20.8384 | -100.0647
20.8421 | -100.0693
20.8458 | -100.0759
20.8494 | -100.0835
20.8531 | -100.0895
20.8568 | -100.0973
20.8605 | -100.1049
20.8641 | -100.1122
20.8678 | -100.1215
20.8715 | -100.13
20.8752 | -100.1366
20.8788 | -100.1461
20.8825 | -100.1565
20.8862 | -100.167
20.8898 | -100.1796
20.8935 | -100.1928
20.8972 | -100.2078
20.9009 | -100.2228
20.9045 | -100.237
20.9082 | -100.2533
20.9119 | -100.2712
20.9156 | -100.2886
20.9192 | -100.3101
20.9229 | -100.3339
20.9266 | -100.3561
20.9303 | -100.3804
20.9339 | -100.4055
20.9376 | -100.43
20.9413 | -100.4566
20.9449 | -100.4834
20.9486 | -100.5097
20.9523 | -100.5379
20.956 | -100.5662
20.9596 | -100.5938
20.9633 | -100.6242
20.967 | -100.6547
20.9707 | -100.6872
20.9743 | -100.7189
20.978 | -100.7499
20.9817 | -100.7804
20.9854 | -100.8122
20.989 | -100.8433
20.9927 | -100.8735
20.9964 | -100.9053
21 | -100.9382
21.0037 | -100.9681
21.0074 | -100.9955
21.0111 | -101.0235
21.0147 | -101.0532
21.0184 | -101.0821
21.0221 | -101.1107
21.0258 | -101.1409
21.0294 | -101.1658
21.0331 | -101.1911
21.0368 | -101.2176
21.0405 | -101.2431
21.0441 | -101.2664
21.0478 | -101.2863
21.0515 | -101.3049
21.0551 | -101.3232
21.0588 | -101.3393
21.0625 | -101.354
21.0662 | -101.3702
21.0698 | -101.3818
21.0735 | -101.3926
21.0772 | -101.3992
21.0809 | -101.4039
21.0845 | -101.4083
21.0882 | -101.4123
21.0919 | -101.4149
21.0956 | -101.419
21.0992 | -101.4218
21.1029 | -101.4197
21.1066 | -101.42
21.1102 | -101.4211
21.1139 | -101.4203
21.1176 | -101.4168
21.1213 | -101.4145
21.1249 | -101.413
21.1286 | -101.4128
21.1323 | -101.4129
21.136 | -101.4106
21.1396 | -101.4109
21.1433 | -101.4109
21.147 | -101.4114
21.1506 | -101.4156
21.1543 | -101.4194
21.158 | -101.4223
21.1617 | -101.4258
21.1653 | -101.4313
21.169 | -101.4394
21.1727 | -101.4487
21.1764 | -101.4577
21.18 | -101.4671
21.1837 | -101.4777
21.1874 | -101.4883
21.1911 | -101.4995
21.1947 | -101.5123
21.1984 | -101.5248
21.2021 | -101.5367
21.2057 | -101.5505
21.2094 | -101.5641
21.2131 | -101.5781
21.2168 | -101.5945
21.2204 | -101.6116
21.2241 | -101.6293
21.2278 | -101.6452
21.2315 | -101.6606
21.2351 | -101.6771
21.2388 | -101.6928
21.2425 | -101.7072
21.2462 | -101.7222
21.2498 | -101.7369
21.2535 | -101.7493
21.2572 | -101.7621
21.2608 | -101.7727
21.2645 | -101.7818
21.2682 | -101.7916
21.2719 | -101.8024
21.2755 | -101.8094
21.2792 | -101.8161
21.2829 | -101.8219
21.2866 | -101.8259
21.2902 | -101.8306
21.2939 | -101.8358
21.2976 | -101.8399
21.3013 | -101.8408
21.3049 | -101.8408
21.3086 | -101.8385
21.3123 | -101.8356
21.3159 | -101.8324
21.3196 | -101.8308
21.3233 | -101.828
21.327 | -101.8232
21.3306 | -101.8178
21.3343 | -101.809
21.338 | -101.7991
21.3417 | -101.7873
21.3453 | -101.7777
21.349 | -101.7681
21.3527 | -101.7565
21.3564 | -101.7456
21.36 | -101.7341
21.3637 | -101.7213
21.3674 | -101.7081
21.371 | -101.6937
21.3747 | -101.679
21.3784 | -101.666
21.3821 | -101.6529
21.3857 | -101.6377
21.3894 | -101.6219
21.3931 | -101.6052
21.3968 | -101.5908
21.4004 | -101.5758
21.4041 | -101.5594
21.4078 | -101.5427
21.4115 | -101.5265
21.4151 | -101.5116
21.4188 | -101.4975
21.4225 | -101.4808
21.4261 | -101.4662
21.4298 | -101.4532
21.4335 | -101.4406
21.4372 | -101.4285
21.4408 | -101.4178
21.4445 | -101.4065
21.4482 | -101.3935
21.4519 | -101.3823
21.4555 | -101.3704
21.4592 | -101.3596
21.4629 | -101.3493
21.4666 | -101.3395
21.4702 | -101.332
21.4739 | -101.3255
21.4776 | -101.3188
21.4812 | -101.3132
21.4849 | -101.3082
21.4886 | -101.3037
21.4923 | -101.3014
21.4959 | -101.3022
21.4996 | -101.3015
21.5033 | -101.3015
21.507 | -101.3021
21.5106 | -101.3023
21.5143 | -101.3042
21.518 | -101.3077
21.5216 | -101.3115
21.5253 | -101.3154
21.529 | -101.3183
21.5327 | -101.3241
21.5363 | -101.3324
21.54 | -101.3418
21.5437 | -101.3515
21.5474 | -101.3613
21.551 | -101.3723
21.5547 | -101.3806
21.5584 | -101.389
21.5621 | -101.3988
21.5657 | -101.4126
21.5694 | -101.4279
21.5731 | -101.4414
21.5767 | -101.455
21.5804 | -101.4712
21.5841 | -101.4862
21.5878 | -101.5017
21.5914 | -101.5157
21.5951 | -101.528
21.5988 | -101.5411
21.6025 | -101.5538
21.6061 | -101.5682
21.6098 | -101.582
21.6135 | -101.5977
21.6172 | -101.6115
21.6208 | -101.6219
21.6245 | -101.6292
21.6282 | -101.6378
21.6318 | -101.647
21.6355 | -101.656
21.6392 | -101.6645
21.6429 | -101.6719
21.6465 | -101.6807
21.6502 | -101.6872
21.6539 | -101.6931
21.6576 | -101.6981
21.6612 | -101.7004
21.6649 | -101.7036
21.6686 | -101.7079
21.6723 | -101.7107
21.6759 | -101.7116
21.6796 | -101.7112
21.6833 | -101.7092
21.6869 | -101.708
21.6906 | -101.709
21.6943 | -101.7078
21.698 | -101.7067
21.7016 | -101.7037
21.7053 | -101.7025
21.709 | -101.7024
21.7127 | -101.7004
21.7163 | -101.6978
21.72 | -101.6932
21.7237 | -101.6882
21.7274 | -101.6861
21.731 | -101.6863
21.7347 | -101.6854
21.7384 | -101.6852
21.742 | -101.6852
21.7457 | -101.6836
21.7494 | -101.6812
21.7531 | -101.6827
21.7567 | -101.6859
21.7604 | -101.6869
21.7641 | -101.6893
21.7678 | -101.6921
21.7714 | -101.6965
21.7751 | -101.7022
21.7788 | -101.7066
21.7825 | -101.7109
21.7861 | -101.7165
21.7898 | -101.7232
21.7935 | -101.7298
21.7971 | -101.7367
21.8008 | -101.7426
21.8045 | -101.7457
21.8082 | -101.7514
21.8118 | -101.7581
21.8155 | -101.7648
21.8192 | -101.773
21.8229 | -101.7807
21.8265 | -101.7866
21.8302 | -101.7919
21.8339 | -101.7983
21.8376 | -101.8077
21.8412 | -101.8154
21.8449 | -101.8208
21.8486 | -101.825
21.8522 | -101.8307
21.8559 | -101.8374
21.8596 | -101.8421
21.8633 | -101.8493
21.8669 | -101.8542
21.8706 | -101.8563
21.8743 | -101.8563
21.878 | -101.8556
21.8816 | -101.8563
21.8853 | -101.8577
21.889 | -101.8587
21.8926 | -101.8596
21.8963 | -101.859
21.9 | -101.8593
21.9037 | -101.8617
21.9073 | -101.8634
21.911 | -101.8641
21.9147 | -101.8655
21.9184 | -101.866
21.922 | -101.8638
21.9257 | -101.8635
21.9294 | -101.8628
21.9331 | -101.8619
21.9367 | -101.8596
21.9404 | -101.855
21.9441 | -101.8515
21.9477 | -101.8502
21.9514 | -101.8488
21.9551 | -101.8449
21.9588 | -101.8405
21.9624 | -101.8369
21.9661 | -101.8334
21.9698 | -101.8282
21.9735 | -101.824
21.9771 | -101.8201
21.9808 | -101.8154
21.9845 | -101.8121
21.9882 | -101.8089
21.9918 | -101.806
21.9955 | -101.802
21.9992 | -101.7977
22.0028 | -101.7935
22.0065 | -101.7889
22.0102 | -101.7837
22.0139 | -101.7778
22.0175 | -101.7725
22.0212 | -101.7691
22.0249 | -101.7657
22.0286 | -101.7625
22.0322 | -101.7595
22.0359 | -101.7571
22.0396 | -101.7524
22.0433 | -101.748
22.0469 | -101.746
22.0506 | -101.7441
22.0543 | -101.7428
22.0579 | -101.7444
22.0616 | -101.7455
22.0653 | -101.7454
22.069 | -101.7462
22.0726 | -101.7485
22.0763 | -101.7517
22.08 | -101.7559
22.0837 | -101.7604
22.0873 | -101.765
22.091 | -101.7723
22.0947 | -101.7781
22.0984 | -101.7842
22.102 | -101.7914
22.1057 | -101.8004
22.1094 | -101.8119
22.113 | -101.8252
22.1167 | -101.8416
22.1204 | -101.8553
22.1241 | -101.8672
22.1277 | -101.8783
22.1314 | -101.8896
22.1351 | -101.9041
22.1388 | -101.9192
22.1424 | -101.9351
22.1461 | -101.9521
22.1498 | -101.9675
22.1535 | -101.9811
22.1571 | -101.9959
22.1608 | -102.0101
22.1645 | -102.0243
22.1681 | -102.0389
22.1718 | -102.0522
22.1755 | -102.0626
22.1792 | -102.0751
22.1828 | -102.0879
22.1865 | -102.1001
22.1902 | -102.1112
22.1939 | -102.1204
22.1975 | -102.1276
22.2012 | -102.1345
22.2049 | -102.1428
22.2085 | -102.1523
22.2122 | -102.1616
22.2159 | -102.1667
22.2196 | -102.1719
22.2232 | -102.1788
22.2269 | -102.1831
22.2306 | -102.1879
22.2343 | -102.1919
22.2379 | -102.1934
22.2416 | -102.1968
22.2453 | -102.1989
22.249 | -102.199
22.2526 | -102.2022
22.2563 | -102.2009
22.26 | -102.2045
22.2636 | -102.2026
22.2673 | -102.2005
22.271 | -102.2001
22.2747 | -102.2001
22.2783 | -102.1986
22.282 | -102.1988
22.2857 | -102.2009
22.2894 | -102.2006
22.293 | -102.2
22.2967 | -102.1989
22.3004 | -102.1956
22.3041 | -102.1921
22.3077 | -102.1897
22.3114 | -102.1885
22.3151 | -102.1862
22.3187 | -102.1848
22.3224 | -102.1834
22.3261 | -102.1829
22.3298 | -102.182
22.3334 | -102.1823
22.3371 | -102.1828
22.3408 | -102.1828
22.3445 | -102.1843
22.3481 | -102.185
22.3518 | -102.1852
22.3555 | -102.1862
22.3592 | -102.1865
22.3628 | -102.1856
22.3665 | -102.1842
22.3702 | -102.1841
22.3738 | -102.1846
22.3775 | -102.186
22.3812 | -102.1887
22.3849 | -102.1911
22.3885 | -102.1932
22.3922 | -102.1945
22.3959 | -102.1959
22.3996 | -102.1967
22.4032 | -102.1975
22.4069 | -102.1988
22.4106 | -102.2039
22.4143 | -102.2087
22.4179 | -102.2135
22.4216 | -102.2181
22.4253 | -102.2214
22.4289 | -102.2245
22.4326 | -102.2266
22.4363 | -102.2291
22.44 | -102.2324
22.4436 | -102.2384
22.4473 | -102.2458
22.451 | -102.2516
22.4547 | -102.2571
22.4583 | -102.2616
22.462 | -102.2666
22.4657 | -102.2725
22.4694 | -102.2799
22.473 | -102.2859
22.4767 | -102.2909
22.4804 | -102.2947
22.484 | -102.2995
22.4877 | -102.3064
22.4914 | -102.3152
22.4951 | -102.3227
22.4987 | -102.3283
22.5024 | -102.332
22.5061 | -102.3332
22.5098 | -102.3344
22.5134 | -102.3364
22.5171 | -102.339
22.5208 | -102.3398
22.5245 | -102.3382
22.5281 | -102.3358
22.5318 | -102.3302
22.5355 | -102.322
22.5391 | -102.3123
22.5428 | -102.3045
22.5465 | -102.2925
22.5502 | -102.2799
22.5538 | -102.2626
22.5575 | -102.2385
22.5612 | -102.2129
22.5649 | -102.1833
22.5685 | -102.1487
22.5722 | -102.1122
22.5759 | -102.0743
22.5795 | -102.0387
22.5832 | -102.0012
22.5869 | -101.96
22.5906 | -101.9182
22.5942 | -101.8751
22.5979 | -101.8301
22.6016 | -101.7817
22.6053 | -101.7324
22.6089 | -101.6823
22.6126 | -101.6331
22.6163 | -101.5819
22.62 | -101.5302
22.6236 | -101.4781
22.6273 | -101.4252
22.631 | -101.3723
22.6346 | -101.3188
22.6383 | -101.2644
22.642 | -101.2118
22.6457 | -101.1587
22.6493 | -101.1069
22.653 | -101.0554
22.6567 | -101.0054
22.6604 | -100.9561
22.664 | -100.9091
22.6677 | -100.8659
22.6714 | -100.8227
22.6751 | -100.7816
22.6787 | -100.7426
22.6824 | -100.7064
22.6861 | -100.6734
22.6897 | -100.6429
22.6934 | -100.6153
22.6971 | -100.5899
22.7008 | -100.5664
22.7044 | -100.547
22.7081 | -100.5311
22.7118 | -100.5184
22.7155 | -100.5091
22.7191 | -100.5015
22.7228 | -100.4975
22.7265 | -100.4938
22.7302 | -100.4922
22.7338 | -100.4952
22.7375 | -100.5005
22.7412 | -100.5095
22.7448 | -100.5208
22.7485 | -100.5343
22.7522 | -100.5519
22.7559 | -100.5722
22.7595 | -100.5946
22.7632 | -100.6173
22.7669 | -100.6397
22.7706 | -100.6643
22.7742 | -100.6915
22.7779 | -100.7204
22.7816 | -100.7504
22.7853 | -100.7822
22.7889 | -100.8097
22.7926 | -100.8391
22.7963 | -100.869
22.7999 | -100.899
22.8036 | -100.9285
22.8073 | -100.9583
22.811 | -100.9871
22.8146 | -101.0139
22.8183 | -101.0388
22.822 | -101.0633
22.8257 | -101.0876
22.8293 | -101.1101
22.833 | -101.1317
22.8367 | -101.153
22.8404 | -101.1713
22.844 | -101.1882
22.8477 | -101.2068
22.8514 | -101.2248
22.855 | -101.2421
22.8587 | -101.257
22.8624 | -101.271
22.8661 | -101.282
22.8697 | -101.2927
22.8734 | -101.3003
22.8771 | -101.3113
22.8808 | -101.3242
22.8844 | -101.3348
22.8881 | -101.3449
22.8918 | -101.3533
22.8955 | -101.3576
22.8991 | -101.3582
22.9028 | -101.352
22.9065 | -101.349
22.9101 | -101.3453
22.9138 | -101.3408
22.9175 | -101.3364
22.9212 | -101.3324
22.9248 | -101.3295
22.9285 | -101.3247
22.9322 | -101.3185
22.9359 | -101.3104
22.9395 | -101.3037
22.9432 | -101.2976
22.9469 | -101.2879
22.9505 | -101.2799
22.9542 | -101.2721
22.9579 | -101.2647
22.9616 | -101.2591
22.9652 | -101.2523
22.9689 | -101.2449
22.9726 | -101.2361
22.9763 | -101.225
22.9799 | -101.2142
22.9836 | -101.2048
22.9873 | -101.1959
22.991 | -101.1871
22.9946 | -101.1787
22.9983 | -101.1691
23.002 | -101.1583
23.0056 | -101.1483
23.0093 | -101.1395
23.013 | -101.1297
23.0167 | -101.1191
23.0203 | -101.1093
23.024 | -101.1005
23.0277 | -101.0903
23.0314 | -101.0804
23.035 | -101.0702
23.0387 | -101.0604
23.0424 | -101.0518
23.0461 | -101.0435
23.0497 | -101.0373
23.0534 | -101.0296
23.0571 | -101.0217
23.0607 | -101.015
23.0644 | -101.0057
23.0681 | -100.998
23.0718 | -100.993
23.0754 | -100.9866
23.0791 | -100.9792
23.0828 | -100.9715
23.0865 | -100.9642
23.0901 | -100.9596
23.0938 | -100.9548
23.0975 | -100.9513
23.1012 | -100.9489
23.1048 | -100.9453
23.1085 | -100.9417
23.1122 | -100.939
23.1158 | -100.9379
23.1195 | -100.9357
23.1232 | -100.9352
23.1269 | -100.9343
23.1305 | -100.932
23.1342 | -100.928
23.1379 | -100.9265
23.1416 | -100.925
23.1452 | -100.9222
23.1489 | -100.92
23.1526 | -100.9195
23.1563 | -100.9177
23.1599 | -100.9155
23.1636 | -100.9144
23.1673 | -100.9148
23.1709 | -100.9147
23.1746 | -100.9104
23.1783 | -100.9051
23.182 | -100.9017
23.1856 | -100.8982
23.1893 | -100.8936
23.193 | -100.8886
23.1967 | -100.8835
23.2003 | -100.8821
23.204 | -100.8778
23.2077 | -100.8732
23.2113 | -100.8698
23.215 | -100.865
23.2187 | -100.8622
23.2224 | -100.8595
23.226 | -100.8562
23.2297 | -100.8553
23.2334 | -100.8513
23.2371 | -100.8474
23.2407 | -100.8424
23.2444 | -100.838
23.2481 | -100.8325
23.2518 | -100.8275
23.2554 | -100.8256
23.2591 | -100.8262
23.2628 | -100.8245
23.2664 | -100.8213
23.2701 | -100.8199
23.2738 | -100.8217
23.2775 | -100.8191
23.2811 | -100.8152
23.2848 | -100.8119
23.2885 | -100.8081
23.2922 | -100.8072
23.2958 | -100.8054
23.2995 | -100.8041
23.3032 | -100.8072
23.3069 | -100.8096
23.3105 | -100.8109
23.3142 | -100.8135
23.3179 | -100.8125
23.3215 | -100.8103
23.3252 | -100.8091
23.3289 | -100.809
23.3326 | -100.8104
23.3362 | -100.8098
23.3399 | -100.8083
23.3436 | -100.8092
23.3473 | -100.8116
23.3509 | -100.8118
23.3546 | -100.811
23.3583 | -100.8125
23.362 | -100.8123
23.3656 | -100.812
23.3693 | -100.8108
23.373 | -100.8123
23.3766 | -100.8135
23.3803 | -100.814
23.384 | -100.815
23.3877 | -100.8137
23.3913 | -100.8114
23.395 | -100.8108
23.3987 | -100.809
23.4024 | -100.8066
23.406 | -100.8043
23.4097 | -100.8029
23.4134 | -100.8026
23.4171 | -100.8013
23.4207 | -100.7981
23.4244 | -100.7957
23.4281 | -100.7935
23.4317 | -100.7911
23.4354 | -100.7894
23.4391 | -100.7881
23.4428 | -100.786
23.4464 | -100.7844
23.4501 | -100.7831
23.4538 | -100.781
23.4575 | -100.7824
23.4611 | -100.7848
23.4648 | -100.786
23.4685 | -100.7861
23.4722 | -100.7876
23.4758 | -100.7902
23.4795 | -100.7933
23.4832 | -100.7959
23.4868 | -100.8007
23.4905 | -100.8047
23.4942 | -100.8065
23.4979 | -100.8101
23.5015 | -100.8149
23.5052 | -100.8214
23.5089 | -100.8292
23.5126 | -100.8374
23.5162 | -100.8473
23.5199 | -100.8571
23.5236 | -100.8677
23.5273 | -100.878
23.5309 | -100.8857
23.5346 | -100.8942
23.5383 | -100.9029
23.5419 | -100.9113
23.5456 | -100.9226
23.5493 | -100.9332
23.553 | -100.9448
23.5566 | -100.9563
23.5603 | -100.9669
23.564 | -100.9784
23.5677 | -100.9912
23.5713 | -101.003
23.575 | -101.0136
23.5787 | -101.0251
23.5824 | -101.0355
23.586 | -101.0451
23.5897 | -101.054
23.5934 | -101.0659
23.597 | -101.0781
23.6007 | -101.0865
23.6044 | -101.0957
23.6081 | -101.108
23.6117 | -101.1201
23.6154 | -101.1311
23.6191 | -101.1397
23.6228 | -101.1463
23.6264 | -101.1564
23.6301 | -101.1651
23.6338 | -101.1719
23.6374 | -101.1783
23.6411 | -101.1836
23.6448 | -101.1888
23.6485 | -101.1945
23.6521 | -101.2
23.6558 | -101.2055
23.6595 | -101.2093
23.6632 | -101.2109
23.6668 | -101.2112
23.6705 | -101.2117
23.6742 | -101.2117
23.6779 | -101.21
23.6815 | -101.2074
23.6852 | -101.2071
23.6889 | -101.2061
23.6925 | -101.2006
23.6962 | -101.1975
23.6999 | -101.195
23.7036 | -101.1922
23.7072 | -101.1864
23.7109 | -101.1788
23.7146 | -101.1726
23.7183 | -101.1677
23.7219 | -101.1603
23.7256 | -101.153
23.7293 | -101.1467
23.733 | -101.1415
23.7366 | -101.1366
23.7403 | -101.1323
23.744 | -101.1275
23.7476 | -101.12
23.7513 | -101.1132
23.755 | -101.1057
23.7587 | -101.0948
23.7623 | -101.0846
23.766 | -101.0748
23.7697 | -101.0669
23.7734 | -101.0604
23.777 | -101.053
23.7807 | -101.0449
23.7844 | -101.0379
23.7881 | -101.0291
23.7917 | -101.0189
23.7954 | -101.0112
23.7991 | -101.0059
23.8027 | -100.9987
23.8064 | -100.9915
23.8101 | -100.9858
23.8138 | -100.977
23.8174 | -100.9674
23.8211 | -100.9609
23.8248 | -100.9552
23.8285 | -100.9493
23.8321 | -100.9414
23.8358 | -100.9338
23.8395 | -100.9277
23.8432 | -100.9224
23.8468 | -100.915
23.8505 | -100.9084
23.8542 | -100.8992
23.8578 | -100.8911
23.8615 | -100.8823
23.8652 | -100.8762
23.8689 | -100.8714
23.8725 | -100.8659
23.8762 | -100.8586
23.8799 | -100.8522
23.8836 | -100.8453
23.8872 | -100.8383
23.8909 | -100.8326
23.8946 | -100.8279
23.8983 | -100.8225
23.9019 | -100.8167
23.9056 | -100.8113
23.9093 | -100.8067
23.9129 | -100.8018
23.9166 | -100.797
23.9203 | -100.7922
23.924 | -100.7854
23.9276 | -100.7801
23.9313 | -100.7742
23.935 | -100.7684
23.9387 | -100.7645
23.9423 | -100.7597
23.946 | -100.7532
23.9497 | -100.7481
23.9534 | -100.7442
23.957 | -100.7423
23.9607 | -100.7401
23.9644 | -100.737
23.968 | -100.7354
23.9717 | -100.7325
23.9754 | -100.7292
23.9791 | -100.7272
23.9827 | -100.7243
23.9864 | -100.7222
23.9901 | -100.7176
23.9938 | -100.7146
23.9974 | -100.7154
24.0011 | -100.7161
24.0048 | -100.7151
24.0084 | -100.7149
24.0121 | -100.7156
24.0158 | -100.7153
24.0195 | -100.7139
24.0231 | -100.7144
24.0268 | -100.7159
24.0305 | -100.7208
24.0342 | -100.7238
24.0378 | -100.7241
24.0415 | -100.7262
24.0452 | -100.7287
24.0489 | -100.7303
24.0525 | -100.7315
24.0562 | -100.7326
24.0599 | -100.7332
24.0635 | -100.7346
24.0672 | -100.7363
24.0709 | -100.7367
24.0746 | -100.7348
24.0782 | -100.7336
24.0819 | -100.7329
24.0856 | -100.7323
24.0893 | -100.7319
24.0929 | -100.7328
24.0966 | -100.734
24.1003 | -100.733
24.104 | -100.7306
24.1076 | -100.726
24.1113 | -100.7198
24.115 | -100.713
24.1186 | -100.7069
24.1223 | -100.7006
24.126 | -100.6941
24.1297 | -100.6849
24.1333 | -100.6744
24.137 | -100.6607
24.1407 | -100.6495
24.1444 | -100.6396
24.148 | -100.628
24.1517 | -100.6187
24.1554 | -100.61
24.1591 | -100.598
24.1627 | -100.5837
24.1664 | -100.57
24.1701 | -100.5566
24.1737 | -100.5408
24.1774 | -100.5258
24.1811 | -100.5097
24.1848 | -100.4964
24.1884 | -100.4839
24.1921 | -100.4671
24.1958 | -100.4512
24.1995 | -100.4364
24.2031 | -100.4235
24.2068 | -100.4074
24.2105 | -100.3894
24.2142 | -100.3744
24.2178 | -100.36
24.2215 | -100.3456
24.2252 | -100.3305
24.2288 | -100.3154
24.2325 | -100.3012
24.2362 | -100.2875
24.2399 | -100.2733
24.2435 | -100.2589
24.2472 | -100.2448
24.2509 | -100.232
24.2546 | -100.2196
24.2582 | -100.2082
24.2619 | -100.1978
24.2656 | -100.1878
24.2692 | -100.1778
24.2729 | -100.1681
24.2766 | -100.1587
24.2803 | -100.1473
24.2839 | -100.1328
24.2876 | -100.1203
24.2913 | -100.1066
24.295 | -100.096
24.2986 | -100.0872
24.3023 | -100.0786
24.306 | -100.071
24.3097 | -100.0643
24.3133 | -100.0577
24.317 | -100.0527
24.3207 | -100.0487
24.3243 | -100.0464
24.328 | -100.0423
24.3317 | -100.0411
24.3354 | -100.0383
24.339 | -100.0331
24.3427 | -100.0296
24.3464 | -100.0276
24.3501 | -100.0254
24.3537 | -100.022
24.3574 | -100.0206
24.3611 | -100.0217
24.3648 | -100.0199
24.3684 | -100.0196
24.3721 | -100.0191
24.3758 | -100.0257
24.3794 | -100.0282
24.3831 | -100.0316
24.3868 | -100.0337
24.3905 | -100.0374
24.3941 | -100.0428
24.3978 | -100.0492
24.4015 | -100.056
24.4052 | -100.0646
24.4088 | -100.0733
24.4125 | -100.0812
24.4162 | -100.0904
24.4199 | -100.1019
24.4235 | -100.1151
24.4272 | -100.1289
24.4309 | -100.1449
24.4345 | -100.1617
24.4382 | -100.1786
24.4419 | -100.1968
24.4456 | -100.2157
24.4492 | -100.2365
24.4529 | -100.2582
24.4566 | -100.2821
24.4603 | -100.3093
24.4639 | -100.3363
24.4676 | -100.3623
24.4713 | -100.3908
24.475 | -100.4213
24.4786 | -100.4515
24.4823 | -100.4847
24.486 | -100.5181
24.4896 | -100.5504
24.4933 | -100.584
24.497 | -100.6181
24.5007 | -100.6525
24.5043 | -100.6874
24.508 | -100.7247
24.5117 | -100.7625
24.5154 | -100.8
24.519 | -100.839
24.5227 | -100.8754
24.5264 | -100.9125
24.5301 | -100.9486
24.5337 | -100.9853
24.5374 | -101.0245
24.5411 | -101.0624
24.5447 | -101.0963
24.5484 | -101.1283
24.5521 | -101.1626
24.5558 | -101.1946
24.5594 | -101.2221
24.5631 | -101.2454
24.5668 | -101.2644
24.5705 | -101.2821
24.5741 | -101.3039
24.5778 | -101.3213
24.5815 | -101.3334
24.5852 | -101.3378
24.5888 | -101.2991
24.5925 | -101.2178
24.5962 | -101.1347
24.5998 | -101.0857
24.6035 | -101.0467
24.6072 | -101.0058
24.6109 | -100.9588
24.6145 | -100.9045
24.6182 | -100.8432
24.6219 | -100.7793
24.6256 | -100.7141
24.6292 | -100.6485
24.6329 | -100.5768
24.6366 | -100.5006
24.6403 | -100.4167
24.6439 | -100.327
24.6476 | -100.2353
24.6513 | -100.1396
24.6549 | -100.0406
24.6586 | -99.9402
24.6623 | -99.8369
24.666 | -99.7308
24.6696 | -99.6233
24.6733 | -99.5179
24.677 | -99.4129
24.6807 | -99.3088
24.6843 | -99.2061
24.688 | -99.1025
24.6917 | -98.9995
24.6953 | -98.899
24.699 | -98.8011
24.7027 | -98.7069
24.7064 | -98.6152
24.71 | -98.5282
24.7137 | -98.4446
24.7174 | -98.3647
24.7211 | -98.2895
24.7247 | -98.2192
24.7284 | -98.1537
24.7321 | -98.0942
24.7358 | -98.0403
24.7394 | -97.9908
24.7431 | -97.9463
24.7468 | -97.9075
24.7504 | -97.8745
24.7541 | -97.8487
24.7578 | -97.8284
24.7615 | -97.8148
24.7651 | -97.8068
24.7688 | -97.8049
24.7725 | -97.8086
24.7762 | -97.8186
24.7798 | -97.8344
24.7835 | -97.8568
24.7872 | -97.8853
24.7909 | -97.9188
24.7945 | -97.9581
24.7982 | -98.0027
24.8019 | -98.053
24.8055 | -98.108
24.8092 | -98.1678
24.8129 | -98.2326
24.8166 | -98.3028
24.8202 | -98.3772
24.8239 | -98.4562
24.8276 | -98.5385
24.8313 | -98.6242
24.8349 | -98.7123
24.8386 | -98.8035
24.8423 | -98.8976
24.846 | -98.9935
24.8496 | -99.091
24.8533 | -99.191
24.857 | -99.2919
24.8606 | -99.3937
24.8643 | -99.4965
24.868 | -99.5997
24.8717 | -99.7027
24.8753 | -99.8041
24.879 | -99.905
24.8827 | -100.0038
24.8864 | -100.1011
24.89 | -100.1961
24.8937 | -100.2892
24.8974 | -100.3802
24.9011 | -100.4682
24.9047 | -100.5146
24.9084 | -100.5976
24.9121 | -100.678
24.9157 | -100.7594
24.9194 | -100.8375
24.9231 | -100.9105
24.9268 | -100.9756
24.9304 | -101.0392
24.9341 | -101.1014
24.9378 | -101.1663
24.9415 | -101.261
24.9451 | -101.3787
24.9488 | -101.4791
24.9525 | -101.5285
24.9562 | -101.5556
24.9598 | -101.578
24.9635 | -101.6002
24.9672 | -101.622
24.9708 | -101.6449
24.9745 | -101.6635
24.9782 | -101.6767
24.9819 | -101.6864
24.9855 | -101.6958
24.9892 | -101.7031
24.9929 | -101.7088
24.9966 | -101.716
25.0002 | -101.7215
25.0039 | -101.7268
25.0076 | -101.7317
25.0112 | -101.7376
25.0149 | -101.744
25.0186 | -101.7496
25.0223 | -101.7545
25.0259 | -101.7569
25.0296 | -101.7585
25.0333 | -101.7603
25.037 | -101.7608
25.0406 | -101.766
25.0443 | -101.7695
25.048 | -101.7704
25.0517 | -101.7733
25.0553 | -101.7748
25.059 | -101.773
25.0627 | -101.7722
25.0663 | -101.7726
25.07 | -101.7733
25.0737 | -101.7757
25.0774 | -101.7766
25.081 | -101.7764
25.0847 | -101.7769
25.0884 | -101.7782
25.0921 | -101.7785
25.0957 | -101.7818
25.0994 | -101.7839
25.1031 | -101.7852
25.1068 | -101.7855
25.1104 | -101.7851
25.1141 | -101.785
25.1178 | -101.7856
25.1214 | -101.7856
25.1251 | -101.7846
25.1288 | -101.7848
25.1325 | -101.7864
25.1361 | -101.7843
25.1398 | -101.784
25.1435 | -101.7867
25.1472 | -101.7882
25.1508 | -101.7889
25.1545 | -101.7894
25.1582 | -101.7878
25.1619 | -101.7852
25.1655 | -101.7814
25.1692 | -101.7776
25.1729 | -101.775
25.1765 | -101.7721
25.1802 | -101.7692
25.1839 | -101.7651
25.1876 | -101.7596
25.1912 | -101.755
25.1949 | -101.7495
25.1986 | -101.746
25.2023 | -101.7403
25.2059 | -101.7314
25.2096 | -101.7231
25.2133 | -101.7149
25.217 | -101.707
25.2206 | -101.6987
25.2243 | -101.6913
25.228 | -101.6834
25.2316 | -101.673
25.2353 | -101.6644
25.239 | -101.6565
25.2427 | -101.6472
25.2463 | -101.6384
25.25 | -101.6297
25.2537 | -101.6225
25.2574 | -101.6162
25.261 | -101.6089
25.2647 | -101.6011
25.2684 | -101.593
25.2721 | -101.5825
25.2757 | -101.5731
25.2794 | -101.5651
25.2831 | -101.5561
25.2867 | -101.5474
25.2904 | -101.5402
25.2941 | -101.5322
25.2978 | -101.5242
25.3014 | -101.5169
25.3051 | -101.5079
25.3088 | -101.5005
25.3125 | -101.4936
25.3161 | -101.4854
25.3198 | -101.4776
25.3235 | -101.4686
25.3271 | -101.4596
25.3308 | -101.45
25.3345 | -101.4419
25.3382 | -101.435
25.3418 | -101.4265
25.3455 | -101.4196
25.3492 | -101.4144
25.3529 | -101.4065
25.3565 | -101.4
25.3602 | -101.3941
25.3639 | -101.3885
25.3676 | -101.3825
25.3712 | -101.3768
25.3749 | -101.3713
25.3786 | -101.3645
25.3822 | -101.3586
25.3859 | -101.3519
25.3896 | -101.3452
25.3933 | -101.3378
25.3969 | -101.3276
25.4006 | -101.3207
25.4043 | -101.3156
25.408 | -101.3081
25.4116 | -101.297
25.4153 | -101.2897
25.419 | -101.2806
25.4227 | -101.2683
25.4263 | -101.2556
25.43 | -101.2439
25.4337 | -101.2301
25.4373 | -101.2194
25.441 | -101.2089
25.4447 | -101.1994
25.4484 | -101.1883
25.452 | -101.1765
25.4557 | -101.165
25.4594 | -101.1538
25.4631 | -101.1428
25.4667 | -101.1325
25.4704 | -101.1225
25.4741 | -101.1109
25.4778 | -101.0999
25.4814 | -101.0887
25.4851 | -101.0754
25.4888 | -101.0637
25.4924 | -101.0528
25.4961 | -101.0434
25.4998 | -101.0302
25.5035 | -101.0171
25.5071 | -101.0045
25.5108 | -100.992
25.5145 | -100.9832
25.5182 | -100.9764
25.5218 | -100.968
25.5255 | -100.9574
25.5292 | -100.948
25.5329 | -100.9389
25.5365 | -100.9306
25.5402 | -100.9235
25.5439 | -100.9176
25.5475 | -100.9143
25.5512 | -100.9081
25.5549 | -100.9008
25.5586 | -100.8934
25.5622 | -100.8878
25.5659 | -100.8839
25.5696 | -100.8802
25.5733 | -100.8756
25.5769 | -100.8711
25.5806 | -100.8678
25.5843 | -100.8648
25.588 | -100.8587
25.5916 | -100.8531
25.5953 | -100.8484
25.599 | -100.8444
25.6026 | -100.8425
25.6063 | -100.8401
25.61 | -100.8352
25.6137 | -100.8304
25.6173 | -100.8264
25.621 | -100.8222
25.6247 | -100.8192
25.6284 | -100.8165
25.632 | -100.8125
25.6357 | -100.809
25.6394 | -100.8072
25.6431 | -100.8053
25.6467 | -100.8023
25.6504 | -100.7996
25.6541 | -100.7963
25.6577 | -100.7927
25.6614 | -100.7906
25.6651 | -100.7885
25.6688 | -100.7837
25.6724 | -100.7803
25.6761 | -100.7773
25.6798 | -100.7723
25.6835 | -100.771
25.6871 | -100.7703
25.6908 | -100.7683
25.6945 | -100.7662
25.6981 | -100.7637
25.7018 | -100.7611
25.7055 | -100.7595
25.7092 | -100.7574
25.7128 | -100.7561
25.7165 | -100.7551
25.7202 | -100.7541
25.7239 | -100.7545
25.7275 | -100.7536
25.7312 | -100.7526
25.7349 | -100.751
25.7386 | -100.7506
25.7422 | -100.7499
25.7459 | -100.7486
25.7496 | -100.748
25.7532 | -100.7489
25.7569 | -100.7495
25.7606 | -100.7493
25.7643 | -100.7498
25.7679 | -100.7489
25.7716 | -100.7501
25.7753 | -100.7518
25.779 | -100.7533
25.7826 | -100.7558
25.7863 | -100.759
25.79 | -100.7599
25.7937 | -100.7611
25.7973 | -100.7617
25.801 | -100.7606
25.8047 | -100.7593
25.8083 | -100.7587
25.812 | -100.7603
25.8157 | -100.7637
25.8194 | -100.7656
25.823 | -100.7652
25.8267 | -100.7625
25.8304 | -100.7622
25.8341 | -100.7613
25.8377 | -100.7598
25.8414 | -100.7606
25.8451 | -100.7613
25.8488 | -100.765
25.8524 | -100.7673
25.8561 | -100.7686
25.8598 | -100.7692
25.8634 | -100.7729
25.8671 | -100.7752
25.8708 | -100.7738
25.8745 | -100.7752
25.8781 | -100.7742
25.8818 | -100.7728
25.8855 | -100.7737
25.8892 | -100.7777
25.8928 | -100.7803
25.8965 | -100.7823
25.9002 | -100.7825
25.9039 | -100.7802
25.9075 | -100.7772
25.9112 | -100.7757
25.9149 | -100.7773
25.9185 | -100.78
25.9222 | -100.7813
25.9259 | -100.7837
25.9296 | -100.7857
25.9332 | -100.7874
25.9369 | -100.7893
25.9406 | -100.7903
25.9443 | -100.7928
25.9479 | -100.7936
25.9516 | -100.7985
25.9553 | -100.805
25.959 | -100.8094
25.9626 | -100.8128
25.9663 | -100.8212
25.97 | -100.8253
25.9736 | -100.8298
25.9773 | -100.8351
25.981 | -100.8394
25.9847 | -100.8439
25.9883 | -100.8493
25.992 | -100.854
25.9957 | -100.8595
25.9994 | -100.8673
26.003 | -100.8745
26.0067 | -100.88
26.0104 | -100.8855
26.0141 | -100.8899
26.0177 | -100.8958
26.0214 | -100.9041
26.0251 | -100.9132
26.0287 | -100.9205
26.0324 | -100.9289
26.0361 | -100.9386
26.0398 | -100.9474
26.0434 | -100.9566
26.0471 | -100.967
26.0508 | -100.9733
26.0545 | -100.9804
26.0581 | -100.9899
26.0618 | -100.9985
26.0655 | -101.005
26.0691 | -101.0113
26.0728 | -101.0166
26.0765 | -101.023
26.0802 | -101.0279
26.0838 | -101.031
26.0875 | -101.035
26.0912 | -101.0401
26.0949 | -101.0442
26.0985 | -101.0461
26.1022 | -101.0482
26.1059 | -101.0515
26.1096 | -101.0558
26.1132 | -101.0608
26.1169 | -101.0644
26.1206 | -101.0659
26.1242 | -101.068
26.1279 | -101.0696
26.1316 | -101.0702
26.1353 | -101.0725
26.1389 | -101.0727
26.1426 | -101.0714
26.1463 | -101.0705
26.15 | -101.0701
26.1536 | -101.0698
26.1573 | -101.0689
26.161 | -101.0676
26.1647 | -101.0656
26.1683 | -101.0652
26.172 | -101.0632
26.1757 | -101.059
26.1793 | -101.0562
26.183 | -101.0539
26.1867 | -101.0517
26.1904 | -101.0507
26.194 | -101.053
26.1977 | -101.0523
26.2014 | -101.0524
26.2051 | -101.054
26.2087 | -101.0531
26.2124 | -101.0531
26.2161 | -101.0525
26.2198 | -101.0503
26.2234 | -101.0504
26.2271 | -101.0512
26.2308 | -101.05
26.2344 | -101.0512
26.2381 | -101.0553
26.2418 | -101.0577
26.2455 | -101.0613
26.2491 | -101.0672
26.2528 | -101.0729
26.2565 | -101.0777
26.2602 | -101.0832
26.2638 | -101.0892
26.2675 | -101.0986
26.2712 | -101.1106
26.2749 | -101.1223
26.2785 | -101.1318
26.2822 | -101.1411
26.2859 | -101.1498
26.2895 | -101.1587
26.2932 | -101.1678
26.2969 | -101.1766
26.3006 | -101.1867
26.3042 | -101.1977
26.3079 | -101.2098
26.3116 | -101.2216
26.3153 | -101.2348
26.3189 | -101.2468
26.3226 | -101.2571
26.3263 | -101.2678
26.33 | -101.2794
26.3336 | -101.2908
26.3373 | -101.3029
26.341 | -101.3148
26.3446 | -101.3247
26.3483 | -101.3353
26.352 | -101.3455
26.3557 | -101.3545
26.3593 | -101.3651
26.363 | -101.3757
26.3667 | -101.3837
26.3704 | -101.391
26.374 | -101.3978
26.3777 | -101.404
26.3814 | -101.4108
26.385 | -101.4185
26.3887 | -101.4233
26.3924 | -101.4276
26.3961 | -101.4288
26.3997 | -101.4285
26.4034 | -101.4281
26.4071 | -101.4304
26.4108 | -101.433
26.4144 | -101.4346
26.4181 | -101.4352
26.4218 | -101.4349
26.4255 | -101.4322
26.4291 | -101.4304
26.4328 | -101.4279
26.4365 | -101.4271
26.4401 | -101.4274
26.4438 | -101.4264
26.4475 | -101.4236
26.4512 | -101.4227
26.4548 | -101.421
26.4585 | -101.4162
26.4622 | -101.4091
26.4659 | -101.4041
26.4695 | -101.403
26.4732 | -101.4002
26.4769 | -101.3925
26.4806 | -101.383
26.4842 | -101.375
26.4879 | -101.3669
26.4916 | -101.3592
26.4952 | -101.3535
26.4989 | -101.3473
26.5026 | -101.3408
26.5063 | -101.3336
26.5099 | -101.3233
26.5136 | -101.3158
26.5173 | -101.3102
26.521 | -101.3047
26.5246 | -101.2985
26.5283 | -101.291
26.532 | -101.2836
26.5357 | -101.2785
26.5393 | -101.2745
26.543 | -101.2696
26.5467 | -101.2681
26.5503 | -101.2664
26.554 | -101.2622
26.5577 | -101.2575
26.5614 | -101.2537
26.565 | -101.2502
26.5687 | -101.2467
26.5724 | -101.2444
26.5761 | -101.2424
26.5797 | -101.2403
26.5834 | -101.2393
26.5871 | -101.2394
26.5908 | -101.2409
26.5944 | -101.2442
26.5981 | -101.2464
26.6018 | -101.2503
26.6054 | -101.2554
26.6091 | -101.2571
26.6128 | -101.2588
26.6165 | -101.2626
26.6201 | -101.2652
26.6238 | -101.267
26.6275 | -101.2704
26.6312 | -101.2747
26.6348 | -101.2789
26.6385 | -101.2837
26.6422 | -101.2885
26.6459 | -101.2942
26.6495 | -101.3006
26.6532 | -101.3076
26.6569 | -101.3141
26.6605 | -101.3216
26.6642 | -101.3303
26.6679 | -101.3381
26.6716 | -101.3473
26.6752 | -101.3542
26.6789 | -101.361
26.6826 | -101.3685
26.6863 | -101.3752
26.6899 | -101.3829
26.6936 | -101.3911
26.6973 | -101.4005
26.7009 | -101.4093
26.7046 | -101.4182
26.7083 | -101.4278
26.712 | -101.4378
26.7156 | -101.4477
26.7193 | -101.4587
26.723 | -101.4689
26.7267 | -101.4777
26.7303 | -101.487
26.734 | -101.4927
26.7377 | -101.4967
26.7414 | -101.5053
26.745 | -101.5149
26.7487 | -101.5231
26.7524 | -101.5315
26.756 | -101.5393
26.7597 | -101.5469
26.7634 | -101.5543
26.7671 | -101.5607
26.7707 | -101.5695
26.7744 | -101.5783
26.7781 | -101.5855
26.7818 | -101.5899
26.7854 | -101.5928
26.7891 | -101.5952
26.7928 | -101.5968
26.7965 | -101.5987
26.8001 | -101.601
26.8038 | -101.6008
26.8075 | -101.6008
26.8111 | -101.6004
26.8148 | -101.5991
26.8185 | -101.5977
26.8222 | -101.5962
26.8258 | -101.5943
26.8295 | -101.592
26.8332 | -101.5904
26.8369 | -101.5887
26.8405 | -101.5862
26.8442 | -101.582
26.8479 | -101.5791
26.8516 | -101.5774
26.8552 | -101.574
26.8589 | -101.5648
26.8626 | -101.5545
26.8662 | -101.5478
26.8699 | -101.5404
26.8736 | -101.5331
26.8773 | -101.5268
26.8809 | -101.5203
26.8846 | -101.5133
26.8883 | -101.5099
26.892 | -101.5045
26.8956 | -101.4945
26.8993 | -101.4869
26.903 | -101.4798
26.9067 | -101.4725
26.9103 | -101.4681
26.914 | -101.4628
26.9177 | -101.4553
26.9213 | -101.4512
26.925 | -101.4475
26.9287 | -101.4444
26.9324 | -101.4418
26.936 | -101.4409
26.9397 | -101.4376
26.9434 | -101.4353
26.9471 | -101.4319
26.9507 | -101.4299
26.9544 | -101.4308
26.9581 | -101.432
26.9618 | -101.4326
26.9654 | -101.4341
26.9691 | -101.4368
26.9728 | -101.4395
26.9764 | -101.442
26.9801 | -101.4458
26.9838 | -101.4489
26.9875 | -101.4561
26.9911 | -101.4661
26.9948 | -101.4731
26.9985 | -101.4802
27.0022 | -101.4869
27.0058 | -101.4929
27.0095 | -101.4989
27.0132 | -101.507
27.0169 | -101.5141
27.0205 | -101.5199
27.0242 | -101.5291
27.0279 | -101.5378
27.0315 | -101.5459
27.0352 | -101.5527
27.0389 | -101.5611
27.0426 | -101.5701
27.0462 | -101.5794
27.0499 | -101.5881
27.0536 | -101.5961
27.0573 | -101.6043
27.0609 | -101.6118
27.0646 | -101.6184
27.0683 | -101.6247
27.072 | -101.6307
27.0756 | -101.6378
27.0793 | -101.6432
27.083 | -101.649
27.0866 | -101.6555
27.0903 | -101.6606
27.094 | -101.6663
27.0977 | -101.6727
27.1013 | -101.6792
27.105 | -101.686
27.1087 | -101.6945
27.1124 | -101.7021
27.116 | -101.7103
27.1197 | -101.7188
27.1234 | -101.7237
27.127 | -101.7262
27.1307 | -101.7314
27.1344 | -101.7366
27.1381 | -101.7397
27.1417 | -101.7444
27.1454 | -101.749
27.1491 | -101.751
27.1528 | -101.753
27.1564 | -101.7571
27.1601 | -101.7628
27.1638 | -101.7682
27.1675 | -101.7736
27.1711 | -101.7775
27.1748 | -101.7818
27.1785 | -101.7881
27.1821 | -101.7919
27.1858 | -101.7958
27.1895 | -101.7983
27.1932 | -101.8012
27.1968 | -101.804
27.2005 | -101.8049
27.2042 | -101.8053
27.2079 | -101.8055
27.2115 | -101.807
27.2152 | -101.8076
27.2189 | -101.8069
27.2226 | -101.8036
27.2262 | -101.8005
27.2299 | -101.7975
27.2336 | -101.7963
27.2372 | -101.7949
27.2409 | -101.7952
27.2446 | -101.7957
27.2483 | -101.7941
27.2519 | -101.7915
27.2556 | -101.7874
27.2593 | -101.7838
27.263 | -101.7804
27.2666 | -101.7763
27.2703 | -101.772
27.274 | -101.7684
27.2777 | -101.7662
27.2813 | -101.762
27.285 | -101.7571
27.2887 | -101.7528
27.2923 | -101.747
27.296 | -101.7408
27.2997 | -101.7372
27.3034 | -101.7311
27.307 | -101.7244
27.3107 | -101.7181
27.3144 | -101.7124
27.3181 | -101.707
27.3217 | -101.7049
27.3254 | -101.7025
27.3291 | -101.7011
27.3328 | -101.6979
27.3364 | -101.6945
27.3401 | -101.6922
27.3438 | -101.6899
27.3474 | -101.687
27.3511 | -101.6838
27.3548 | -101.6823
27.3585 | -101.6826
27.3621 | -101.6814
27.3658 | -101.6815
27.3695 | -101.6825
27.3732 | -101.6827
27.3768 | -101.6821
27.3805 | -101.6839
27.3842 | -101.6863
27.3879 | -101.6895
27.3915 | -101.6937
27.3952 | -101.6983
27.3989 | -101.7044
27.4025 | -101.7116
27.4062 | -101.7169
27.4099 | -101.7227
27.4136 | -101.7285
27.4172 | -101.7366
27.4209 | -101.7458
27.4246 | -101.7534
27.4283 | -101.7619
27.4319 | -101.7704
27.4356 | -101.7791
27.4393 | -101.7891
27.4429 | -101.8003
27.4466 | -101.8123
27.4503 | -101.8248
27.454 | -101.8371
27.4576 | -101.8496
27.4613 | -101.8605
27.465 | -101.8726
27.4687 | -101.8846
27.4723 | -101.8956
27.476 | -101.9071
27.4797 | -101.9177
27.4834 | -101.9285
27.487 | -101.9406
27.4907 | -101.9508
27.4944 | -101.9609
27.498 | -101.9723
27.5017 | -101.9854
27.5054 | -101.9988
27.5091 | -102.0108
27.5127 | -102.0248
27.5164 | -102.0369
27.5201 | -102.0475
27.5238 | -102.0603
27.5274 | -102.0724
27.5311 | -102.0849
27.5348 | -102.0995
27.5385 | -102.1133
27.5421 | -102.1253
27.5458 | -102.1358
27.5495 | -102.1453
27.5531 | -102.1554
27.5568 | -102.1648
27.5605 | -102.1762
27.5642 | -102.1874
27.5678 | -102.195
27.5715 | -102.2042
27.5752 | -102.2138
27.5789 | -102.2218
27.5825 | -102.2297
27.5862 | -102.236
27.5899 | -102.2435
27.5936 | -102.2521
27.5972 | -102.26
27.6009 | -102.2643
27.6046 | -102.2692
27.6082 | -102.2737
27.6119 | -102.2798
27.6156 | -102.2856
27.6193 | -102.2903
27.6229 | -102.2913
27.6266 | -102.2935
27.6303 | -102.2953
27.634 | -102.2946
27.6376 | -102.2912
27.6413 | -102.2856
27.645 | -102.2813
27.6487 | -102.2756
27.6523 | -102.2704
27.656 | -102.2649
27.6597 | -102.2571
27.6633 | -102.2487
27.667 | -102.2388
27.6707 | -102.2274
27.6744 | -102.2196
27.678 | -102.2112
27.6817 | -102.201
27.6854 | -102.1906
27.6891 | -102.1786
27.6927 | -102.1656
27.6964 | -102.1516
27.7001 | -102.1377
27.7038 | -102.1225
27.7074 | -102.1062
27.7111 | -102.089
27.7148 | -102.0719
27.7184 | -102.0549
27.7221 | -102.0385
27.7258 | -102.0198
27.7295 | -102.0019
27.7331 | -101.9856
27.7368 | -101.9685
27.7405 | -101.9505
27.7442 | -101.9322
27.7478 | -101.9143
27.7515 | -101.8974
27.7552 | -101.8813
27.7588 | -101.8652
27.7625 | -101.8494
27.7662 | -101.8368
27.7699 | -101.8231
27.7735 | -101.8102
27.7772 | -101.7986
27.7809 | -101.7878
27.7846 | -101.7776
27.7882 | -101.7686
27.7919 | -101.7595
27.7956 | -101.7509
27.7993 | -101.744
27.8029 | -101.7373
27.8066 | -101.7304
27.8103 | -101.7252
27.8139 | -101.7216
27.8176 | -101.7205
27.8213 | -101.7204
27.825 | -101.7223
27.8286 | -101.7237
27.8323 | -101.7246
27.836 | -101.7268
27.8397 | -101.7289
27.8433 | -101.7329
27.847 | -101.7401
27.8507 | -101.7483
27.8544 | -101.7569
27.858 | -101.7669
27.8617 | -101.7764
27.8654 | -101.7876
27.869 | -101.8009
27.8727 | -101.8147
27.8764 | -101.8276
27.8801 | -101.8418
27.8837 | -101.8556
27.8874 | -101.8687
27.8911 | -101.8819
27.8948 | -101.8961
27.8984 | -101.9107
27.9021 | -101.9255
27.9058 | -101.9402
27.9095 | -101.9556
27.9131 | -101.9709
27.9168 | -101.9856
27.9205 | -101.9992
27.9241 | -102.0128
27.9278 | -102.0258
27.9315 | -102.0378
27.9352 | -102.0471
27.9388 | -102.0597
27.9425 | -102.0727
27.9462 | -102.0852
27.9499 | -102.0988
27.9535 | -102.1107
27.9572 | -102.1226
27.9609 | -102.1348
27.9646 | -102.145
27.9682 | -102.1552
27.9719 | -102.1666
27.9756 | -102.1755
27.9792 | -102.1844
27.9829 | -102.1931
27.9866 | -102.1993
27.9903 | -102.2053
27.9939 | -102.2124
27.9976 | -102.2184
28.0013 | -102.2237