## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 232 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 234
Hood | 02-92B/P | 173
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1008.5 | hPa |  | PRC PRO 01
Humidity | 45.1 | %/Hr
Internal temperature of calibration bench | 23.99 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1021 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.7 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.88 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.04 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3064 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 22/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 25/05/2018
Engineer | Anthony HUE | Approval |  | 29/05/2018
CEO | Robert GLEMOT | Approval |  | 29/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4842 | 57.6311
0.012209 | -36.0421 | 38.4303
0.018313 | -34.9724 | 28.2987
0.024417 | -34.5064 | 22.0444
0.030521 | -34.3965 | 18.2881
0.036626 | -34.251 | 15.2248
0.04273 | -34.1545 | 12.8444
0.048834 | -34.1082 | 11.3282
0.054938 | -34.0411 | 10.1993
0.061043 | -33.9781 | 9.1146
0.067147 | -33.968 | 8.2892
0.073251 | -33.9693 | 7.527
0.079355 | -33.9883 | 6.8633
0.08546 | -34.0027 | 6.2663
0.091564 | -33.9975 | 5.8256
0.097668 | -33.9783 | 5.5816
0.10377 | -33.9547 | 5.3371
0.10988 | -33.9364 | 5.0614
0.11598 | -33.9267 | 4.7688
0.12209 | -33.9286 | 4.4842
0.12819 | -33.9392 | 4.1989
0.13429 | -33.9508 | 3.908
0.1404 | -33.9544 | 3.6261
0.1465 | -33.9565 | 3.3655
0.15261 | -33.9535 | 3.1436
0.15871 | -33.9465 | 2.9817
0.16482 | -33.9377 | 2.8812
0.17092 | -33.935 | 2.8458
0.17702 | -33.9324 | 2.7759
0.18313 | -33.9313 | 2.6992
0.18923 | -33.9307 | 2.5902
0.19534 | -33.9293 | 2.4623
0.20144 | -33.9279 | 2.3181
0.20754 | -33.9247 | 2.1958
0.21365 | -33.9236 | 2.08
0.21975 | -33.923 | 1.9815
0.22586 | -33.9239 | 1.8938
0.23196 | -33.9263 | 1.8217
0.23807 | -33.9304 | 1.7595
0.24417 | -33.9345 | 1.7014
0.25027 | -33.9385 | 1.6396
0.25638 | -33.9408 | 1.5781
0.26248 | -33.9426 | 1.5117
0.26859 | -33.9434 | 1.4468
0.27469 | -33.9436 | 1.3757
0.2808 | -33.9446 | 1.3017
0.2869 | -33.9453 | 1.2306
0.293 | -33.9461 | 1.1634
0.29911 | -33.947 | 1.1034
0.30521 | -33.9479 | 1.0443
0.31132 | -33.9486 | 1.0061
0.31742 | -33.9497 | 0.96367
0.32353 | -33.95 | 0.92928
0.32963 | -33.9508 | 0.90096
0.33573 | -33.9517 | 0.88169
0.34184 | -33.953 | 0.86115
0.34794 | -33.9551 | 0.83808
0.35405 | -33.9573 | 0.8147
0.36015 | -33.9597 | 0.79001
0.36626 | -33.9617 | 0.76307
0.37236 | -33.9638 | 0.72658
0.37846 | -33.9659 | 0.68221
0.38457 | -33.9667 | 0.62434
0.39067 | -33.9677 | 0.57739
0.39678 | -33.9682 | 0.52493
0.40288 | -33.9686 | 0.47021
0.40899 | -33.9683 | 0.41765
0.41509 | -33.9675 | 0.36606
0.42119 | -33.9674 | 0.32226
0.4273 | -33.9672 | 0.27545
0.4334 | -33.967 | 0.2287
0.43951 | -33.9672 | 0.18478
0.44561 | -33.968 | 0.14464
0.45172 | -33.9696 | 0.11272
0.45782 | -33.9704 | 0.075283
0.46392 | -33.9716 | 0.038739
0.47003 | -33.9727 | 0.0061275
0.47613 | -33.9735 | -0.028265
0.48224 | -33.974 | -0.063423
0.48834 | -33.973 | -0.098537
0.49445 | -33.9727 | -0.13418
0.50055 | -33.9721 | -0.16648
0.50665 | -33.9713 | -0.20104
0.51276 | -33.9702 | -0.23542
0.51886 | -33.9691 | -0.26968
0.52497 | -33.9677 | -0.29376
0.53107 | -33.9668 | -0.32334
0.53717 | -33.966 | -0.35161
0.54328 | -33.9651 | -0.3743
0.54938 | -33.9648 | -0.39662
0.55549 | -33.9647 | -0.41729
0.56159 | -33.9658 | -0.42835
0.5677 | -33.966 | -0.44252
0.5738 | -33.967 | -0.45848
0.5799 | -33.9678 | -0.47303
0.58601 | -33.9688 | -0.48264
0.59211 | -33.9705 | -0.48837
0.59822 | -33.9718 | -0.49735
0.60432 | -33.9733 | -0.50904
0.61043 | -33.9746 | -0.51653
0.61653 | -33.9758 | -0.52644
0.62263 | -33.9774 | -0.53545
0.62874 | -33.978 | -0.5459
0.63484 | -33.9789 | -0.55486
0.64095 | -33.9793 | -0.56481
0.64705 | -33.9797 | -0.57435
0.65316 | -33.9798 | -0.58667
0.65926 | -33.9799 | -0.60248
0.66536 | -33.9787 | -0.62
0.67147 | -33.9782 | -0.63315
0.67757 | -33.9778 | -0.64742
0.68368 | -33.977 | -0.66645
0.68978 | -33.9761 | -0.68374
0.69589 | -33.9751 | -0.70288
0.70199 | -33.974 | -0.72586
0.70809 | -33.9731 | -0.74517
0.7142 | -33.9723 | -0.76579
0.7203 | -33.9711 | -0.78512
0.72641 | -33.9703 | -0.80781
0.73251 | -33.9696 | -0.82792
0.73862 | -33.9698 | -0.84978
0.74472 | -33.9694 | -0.86987
0.75082 | -33.9691 | -0.89212
0.75693 | -33.9694 | -0.91408
0.76303 | -33.9699 | -0.9362
0.76914 | -33.9718 | -0.95665
0.77524 | -33.9727 | -0.97631
0.78135 | -33.974 | -0.99641
0.78745 | -33.9754 | -1.0164
0.79355 | -33.9766 | -1.0361
0.79966 | -33.9784 | -1.0555
0.80576 | -33.9807 | -1.0751
0.81187 | -33.9824 | -1.0967
0.81797 | -33.9843 | -1.118
0.82408 | -33.9861 | -1.1385
0.83018 | -33.9878 | -1.159
0.83628 | -33.9898 | -1.1774
0.84239 | -33.9909 | -1.198
0.84849 | -33.9924 | -1.2197
0.8546 | -33.9939 | -1.2396
0.8607 | -33.9951 | -1.2603
0.86681 | -33.9962 | -1.2792
0.87291 | -33.9972 | -1.298
0.87901 | -33.9977 | -1.316
0.88512 | -33.9984 | -1.3341
0.89122 | -33.9991 | -1.3523
0.89733 | -33.9995 | -1.3678
0.90343 | -33.9999 | -1.3854
0.90953 | -34 | -1.3996
0.91564 | -34 | -1.4154
0.92174 | -34.0002 | -1.4302
0.92785 | -34.0003 | -1.4447
0.93395 | -34.0003 | -1.4584
0.94006 | -34.0005 | -1.4711
0.94616 | -34.0001 | -1.4827
0.95226 | -34.0002 | -1.4952
0.95837 | -34.0003 | -1.509
0.96447 | -34.0004 | -1.5216
0.97058 | -34.0002 | -1.5349
0.97668 | -34 | -1.5481
0.98279 | -33.9996 | -1.5607
0.98889 | -33.9994 | -1.5712
0.99499 | -33.9993 | -1.5845
1.0011 | -33.9991 | -1.5949
1.0072 | -33.9987 | -1.6064
1.0133 | -33.9983 | -1.6169
1.0194 | -33.9979 | -1.6298
1.0255 | -33.9975 | -1.6434
1.0316 | -33.9968 | -1.6552
1.0377 | -33.9964 | -1.6687
1.0438 | -33.9957 | -1.6813
1.0499 | -33.9949 | -1.6965
1.056 | -33.9941 | -1.7108
1.0621 | -33.9934 | -1.725
1.0682 | -33.9929 | -1.7373
1.0743 | -33.9922 | -1.7527
1.0805 | -33.9916 | -1.7669
1.0866 | -33.9906 | -1.7819
1.0927 | -33.9897 | -1.7963
1.0988 | -33.989 | -1.8111
1.1049 | -33.9881 | -1.8252
1.111 | -33.9872 | -1.8405
1.1171 | -33.9864 | -1.8549
1.1232 | -33.9859 | -1.87
1.1293 | -33.9854 | -1.8839
1.1354 | -33.9851 | -1.8988
1.1415 | -33.9847 | -1.9142
1.1476 | -33.984 | -1.9296
1.1537 | -33.9834 | -1.9442
1.1598 | -33.9835 | -1.9588
1.1659 | -33.9832 | -1.9719
1.172 | -33.9829 | -1.9856
1.1781 | -33.9826 | -1.9994
1.1842 | -33.9824 | -2.0135
1.1903 | -33.9828 | -2.0242
1.1964 | -33.9826 | -2.0374
1.2025 | -33.9827 | -2.0512
1.2086 | -33.9831 | -2.0633
1.2147 | -33.983 | -2.0756
1.2209 | -33.9834 | -2.0865
1.227 | -33.9841 | -2.097
1.2331 | -33.9845 | -2.1088
1.2392 | -33.9849 | -2.1204
1.2453 | -33.9853 | -2.1312
1.2514 | -33.9857 | -2.1408
1.2575 | -33.9863 | -2.1515
1.2636 | -33.9872 | -2.1602
1.2697 | -33.9878 | -2.1692
1.2758 | -33.9885 | -2.1801
1.2819 | -33.9892 | -2.1898
1.288 | -33.9899 | -2.2003
1.2941 | -33.9907 | -2.2085
1.3002 | -33.9914 | -2.2162
1.3063 | -33.992 | -2.225
1.3124 | -33.9926 | -2.2338
1.3185 | -33.9934 | -2.2426
1.3246 | -33.9941 | -2.2514
1.3307 | -33.9949 | -2.2597
1.3368 | -33.9957 | -2.2681
1.3429 | -33.9964 | -2.2768
1.349 | -33.9969 | -2.2859
1.3551 | -33.9976 | -2.295
1.3613 | -33.9981 | -2.303
1.3674 | -33.9988 | -2.3128
1.3735 | -33.9992 | -2.3224
1.3796 | -33.9997 | -2.3312
1.3857 | -34 | -2.3402
1.3918 | -34.0005 | -2.3501
1.3979 | -34.0011 | -2.359
1.404 | -34.0015 | -2.3702
1.4101 | -34.002 | -2.3803
1.4162 | -34.0023 | -2.3906
1.4223 | -34.0028 | -2.4018
1.4284 | -34.0033 | -2.4116
1.4345 | -34.0035 | -2.4229
1.4406 | -34.0038 | -2.4363
1.4467 | -34.0042 | -2.4477
1.4528 | -34.0047 | -2.4588
1.4589 | -34.0052 | -2.4705
1.465 | -34.0054 | -2.4812
1.4711 | -34.0059 | -2.4923
1.4772 | -34.0064 | -2.5055
1.4833 | -34.0068 | -2.5184
1.4894 | -34.0073 | -2.5304
1.4955 | -34.0079 | -2.5422
1.5016 | -34.0084 | -2.5545
1.5078 | -34.0088 | -2.5683
1.5139 | -34.0091 | -2.5812
1.52 | -34.0096 | -2.5936
1.5261 | -34.0099 | -2.6061
1.5322 | -34.0102 | -2.62
1.5383 | -34.0105 | -2.6324
1.5444 | -34.0111 | -2.6449
1.5505 | -34.0115 | -2.6582
1.5566 | -34.012 | -2.6711
1.5627 | -34.0125 | -2.6847
1.5688 | -34.0131 | -2.6982
1.5749 | -34.0134 | -2.7106
1.581 | -34.0137 | -2.7238
1.5871 | -34.0141 | -2.7375
1.5932 | -34.0144 | -2.7502
1.5993 | -34.0148 | -2.7643
1.6054 | -34.0151 | -2.7774
1.6115 | -34.0156 | -2.7914
1.6176 | -34.0158 | -2.8038
1.6237 | -34.016 | -2.8173
1.6298 | -34.0162 | -2.8296
1.6359 | -34.0166 | -2.8422
1.642 | -34.017 | -2.8539
1.6482 | -34.0172 | -2.8668
1.6543 | -34.0175 | -2.8803
1.6604 | -34.0178 | -2.8921
1.6665 | -34.0179 | -2.9044
1.6726 | -34.0181 | -2.9167
1.6787 | -34.0184 | -2.9291
1.6848 | -34.0183 | -2.9414
1.6909 | -34.0185 | -2.9539
1.697 | -34.0187 | -2.9659
1.7031 | -34.0187 | -2.9783
1.7092 | -34.0186 | -2.9902
1.7153 | -34.0188 | -3.002
1.7214 | -34.0185 | -3.014
1.7275 | -34.0184 | -3.0264
1.7336 | -34.0183 | -3.0391
1.7397 | -34.0185 | -3.0513
1.7458 | -34.0184 | -3.0632
1.7519 | -34.0184 | -3.0752
1.758 | -34.0184 | -3.0869
1.7641 | -34.0185 | -3.0983
1.7702 | -34.0184 | -3.111
1.7763 | -34.0184 | -3.1223
1.7824 | -34.0183 | -3.1345
1.7885 | -34.0183 | -3.1468
1.7947 | -34.0183 | -3.1586
1.8008 | -34.0183 | -3.1696
1.8069 | -34.0183 | -3.1809
1.813 | -34.0183 | -3.1928
1.8191 | -34.0183 | -3.2054
1.8252 | -34.0183 | -3.218
1.8313 | -34.0184 | -3.23
1.8374 | -34.0185 | -3.2417
1.8435 | -34.0186 | -3.2534
1.8496 | -34.0186 | -3.2663
1.8557 | -34.0187 | -3.2789
1.8618 | -34.019 | -3.291
1.8679 | -34.019 | -3.3028
1.874 | -34.0191 | -3.3147
1.8801 | -34.0193 | -3.327
1.8862 | -34.0195 | -3.3387
1.8923 | -34.0196 | -3.3507
1.8984 | -34.0197 | -3.3627
1.9045 | -34.0197 | -3.3741
1.9106 | -34.0197 | -3.3862
1.9167 | -34.0198 | -3.3983
1.9228 | -34.0199 | -3.4099
1.9289 | -34.02 | -3.4217
1.9351 | -34.0201 | -3.4324
1.9412 | -34.0203 | -3.4435
1.9473 | -34.0206 | -3.4555
1.9534 | -34.0208 | -3.4668
1.9595 | -34.0209 | -3.4785
1.9656 | -34.0212 | -3.4908
1.9717 | -34.0214 | -3.5024
1.9778 | -34.0217 | -3.5136
1.9839 | -34.0219 | -3.525
1.99 | -34.0221 | -3.536
1.9961 | -34.0221 | -3.548
2.0022 | -34.0224 | -3.5594
2.0083 | -34.0224 | -3.5705
2.0144 | -34.0226 | -3.5818
2.0205 | -34.0229 | -3.5935
2.0266 | -34.0233 | -3.6048
2.0327 | -34.0235 | -3.6164
2.0388 | -34.0239 | -3.6279
2.0449 | -34.0241 | -3.6392
2.051 | -34.0243 | -3.6505
2.0571 | -34.0246 | -3.6618
2.0632 | -34.0248 | -3.6731
2.0693 | -34.025 | -3.6853
2.0754 | -34.0253 | -3.6966
2.0816 | -34.0256 | -3.7087
2.0877 | -34.0257 | -3.7201
2.0938 | -34.0261 | -3.7322
2.0999 | -34.0263 | -3.7438
2.106 | -34.0264 | -3.7554
2.1121 | -34.0266 | -3.7673
2.1182 | -34.0269 | -3.778
2.1243 | -34.0271 | -3.7887
2.1304 | -34.0273 | -3.7998
2.1365 | -34.0276 | -3.8114
2.1426 | -34.0278 | -3.8234
2.1487 | -34.0281 | -3.8345
2.1548 | -34.0285 | -3.8463
2.1609 | -34.029 | -3.8577
2.167 | -34.0294 | -3.8688
2.1731 | -34.0297 | -3.8807
2.1792 | -34.0299 | -3.8917
2.1853 | -34.0302 | -3.9036
2.1914 | -34.0305 | -3.9149
2.1975 | -34.0309 | -3.9264
2.2036 | -34.0311 | -3.9384
2.2097 | -34.0314 | -3.9498
2.2158 | -34.0318 | -3.9604
2.222 | -34.032 | -3.9707
2.2281 | -34.0323 | -3.9821
2.2342 | -34.0326 | -3.9933
2.2403 | -34.0327 | -4.0038
2.2464 | -34.033 | -4.0144
2.2525 | -34.0333 | -4.0257
2.2586 | -34.0336 | -4.0362
2.2647 | -34.0339 | -4.0465
2.2708 | -34.0342 | -4.0576
2.2769 | -34.0346 | -4.0682
2.283 | -34.0347 | -4.0792
2.2891 | -34.035 | -4.09
2.2952 | -34.0351 | -4.1004
2.3013 | -34.0354 | -4.1111
2.3074 | -34.0356 | -4.1214
2.3135 | -34.0358 | -4.1327
2.3196 | -34.036 | -4.1427
2.3257 | -34.0362 | -4.1533
2.3318 | -34.0364 | -4.1642
2.3379 | -34.0366 | -4.1749
2.344 | -34.0367 | -4.1861
2.3501 | -34.0369 | -4.1967
2.3562 | -34.037 | -4.2069
2.3623 | -34.0372 | -4.2171
2.3685 | -34.0373 | -4.2275
2.3746 | -34.0374 | -4.2386
2.3807 | -34.0376 | -4.2489
2.3868 | -34.0378 | -4.2594
2.3929 | -34.038 | -4.2699
2.399 | -34.038 | -4.2802
2.4051 | -34.0383 | -4.2906
2.4112 | -34.0385 | -4.301
2.4173 | -34.0387 | -4.3115
2.4234 | -34.0387 | -4.3222
2.4295 | -34.0389 | -4.3328
2.4356 | -34.039 | -4.3435
2.4417 | -34.0391 | -4.3545
2.4478 | -34.0392 | -4.3652
2.4539 | -34.0394 | -4.3752
2.46 | -34.0395 | -4.3865
2.4661 | -34.0396 | -4.3974
2.4722 | -34.0398 | -4.4074
2.4783 | -34.0398 | -4.4177
2.4844 | -34.0399 | -4.4285
2.4905 | -34.04 | -4.4393
2.4966 | -34.0401 | -4.4499
2.5027 | -34.0402 | -4.4602
2.5089 | -34.0403 | -4.4709
2.515 | -34.0403 | -4.4816
2.5211 | -34.0404 | -4.4921
2.5272 | -34.0405 | -4.5028
2.5333 | -34.0406 | -4.5136
2.5394 | -34.0407 | -4.5242
2.5455 | -34.0408 | -4.5355
2.5516 | -34.041 | -4.5456
2.5577 | -34.0411 | -4.5569
2.5638 | -34.0413 | -4.5685
2.5699 | -34.0413 | -4.5797
2.576 | -34.0413 | -4.5911
2.5821 | -34.0414 | -4.6022
2.5882 | -34.0415 | -4.6132
2.5943 | -34.0416 | -4.6238
2.6004 | -34.0418 | -4.6346
2.6065 | -34.0419 | -4.6448
2.6126 | -34.042 | -4.6551
2.6187 | -34.0421 | -4.6654
2.6248 | -34.0422 | -4.676
2.6309 | -34.0424 | -4.6872
2.637 | -34.0426 | -4.6987
2.6431 | -34.0427 | -4.709
2.6492 | -34.0428 | -4.7202
2.6554 | -34.043 | -4.7309
2.6615 | -34.043 | -4.7417
2.6676 | -34.0433 | -4.7529
2.6737 | -34.0435 | -4.764
2.6798 | -34.0438 | -4.7749
2.6859 | -34.0439 | -4.786
2.692 | -34.0441 | -4.7967
2.6981 | -34.0443 | -4.8076
2.7042 | -34.0444 | -4.8186
2.7103 | -34.0446 | -4.8291
2.7164 | -34.0449 | -4.8399
2.7225 | -34.0451 | -4.8511
2.7286 | -34.0453 | -4.8619
2.7347 | -34.0455 | -4.8723
2.7408 | -34.0458 | -4.8822
2.7469 | -34.046 | -4.8934
2.753 | -34.0462 | -4.9038
2.7591 | -34.0465 | -4.9149
2.7652 | -34.0467 | -4.9262
2.7713 | -34.0469 | -4.9368
2.7774 | -34.0471 | -4.9476
2.7835 | -34.0473 | -4.9584
2.7896 | -34.0475 | -4.9687
2.7958 | -34.0477 | -4.9793
2.8019 | -34.048 | -4.9895
2.808 | -34.0482 | -5.0002
2.8141 | -34.0485 | -5.0112
2.8202 | -34.0488 | -5.0219
2.8263 | -34.0491 | -5.0326
2.8324 | -34.0495 | -5.0438
2.8385 | -34.0498 | -5.0545
2.8446 | -34.0502 | -5.0651
2.8507 | -34.0505 | -5.0758
2.8568 | -34.0509 | -5.0861
2.8629 | -34.0511 | -5.0965
2.869 | -34.0514 | -5.1075
2.8751 | -34.0517 | -5.118
2.8812 | -34.0521 | -5.1292
2.8873 | -34.0524 | -5.14
2.8934 | -34.0527 | -5.1508
2.8995 | -34.0531 | -5.1618
2.9056 | -34.0533 | -5.1728
2.9117 | -34.0536 | -5.1837
2.9178 | -34.054 | -5.1942
2.9239 | -34.0544 | -5.2056
2.93 | -34.0548 | -5.217
2.9361 | -34.0551 | -5.2276
2.9423 | -34.0554 | -5.2382
2.9484 | -34.0558 | -5.2491
2.9545 | -34.0562 | -5.2599
2.9606 | -34.0566 | -5.2712
2.9667 | -34.0569 | -5.2822
2.9728 | -34.0573 | -5.2934
2.9789 | -34.0577 | -5.3043
2.985 | -34.0581 | -5.3155
2.9911 | -34.0585 | -5.3267
2.9972 | -34.0589 | -5.3376
3.0033 | -34.0593 | -5.349
3.0094 | -34.0597 | -5.36
3.0155 | -34.0601 | -5.3708
3.0216 | -34.0605 | -5.3819
3.0277 | -34.0608 | -5.3935
3.0338 | -34.0612 | -5.4047
3.0399 | -34.0616 | -5.4159
3.046 | -34.0619 | -5.4274
3.0521 | -34.0623 | -5.4384
3.0582 | -34.0625 | -5.4503
3.0643 | -34.0629 | -5.4617
3.0704 | -34.0632 | -5.4734
3.0765 | -34.0635 | -5.4844
3.0827 | -34.0638 | -5.4955
3.0888 | -34.0641 | -5.5067
3.0949 | -34.0645 | -5.5179
3.101 | -34.0649 | -5.5293
3.1071 | -34.0652 | -5.5409
3.1132 | -34.0656 | -5.5526
3.1193 | -34.0659 | -5.5639
3.1254 | -34.0663 | -5.5752
3.1315 | -34.0667 | -5.587
3.1376 | -34.0669 | -5.5983
3.1437 | -34.0673 | -5.6098
3.1498 | -34.0676 | -5.6217
3.1559 | -34.0679 | -5.6331
3.162 | -34.0682 | -5.6451
3.1681 | -34.0684 | -5.657
3.1742 | -34.0688 | -5.6685
3.1803 | -34.0691 | -5.68
3.1864 | -34.0694 | -5.6912
3.1925 | -34.0697 | -5.7027
3.1986 | -34.07 | -5.7146
3.2047 | -34.0703 | -5.7266
3.2108 | -34.0705 | -5.7375
3.2169 | -34.0708 | -5.7492
3.223 | -34.0711 | -5.7608
3.2292 | -34.0713 | -5.7726
3.2353 | -34.0715 | -5.7833
3.2414 | -34.0717 | -5.7947
3.2475 | -34.0719 | -5.8059
3.2536 | -34.0722 | -5.8173
3.2597 | -34.0725 | -5.8282
3.2658 | -34.0727 | -5.839
3.2719 | -34.0729 | -5.8499
3.278 | -34.0731 | -5.8614
3.2841 | -34.0734 | -5.8724
3.2902 | -34.0736 | -5.8837
3.2963 | -34.0739 | -5.8951
3.3024 | -34.074 | -5.9061
3.3085 | -34.0743 | -5.9176
3.3146 | -34.0744 | -5.9286
3.3207 | -34.0747 | -5.94
3.3268 | -34.0749 | -5.9514
3.3329 | -34.075 | -5.9621
3.339 | -34.0752 | -5.9727
3.3451 | -34.0754 | -5.9839
3.3512 | -34.0755 | -5.9945
3.3573 | -34.0758 | -6.0056
3.3634 | -34.0759 | -6.0167
3.3696 | -34.0761 | -6.0279
3.3757 | -34.0762 | -6.0386
3.3818 | -34.0765 | -6.0494
3.3879 | -34.0766 | -6.0604
3.394 | -34.0768 | -6.0716
3.4001 | -34.077 | -6.0823
3.4062 | -34.077 | -6.0932
3.4123 | -34.0772 | -6.1037
3.4184 | -34.0774 | -6.1148
3.4245 | -34.0776 | -6.1255
3.4306 | -34.0776 | -6.1361
3.4367 | -34.0778 | -6.1465
3.4428 | -34.078 | -6.1568
3.4489 | -34.0781 | -6.1676
3.455 | -34.0782 | -6.1784
3.4611 | -34.0784 | -6.1895
3.4672 | -34.0785 | -6.2005
3.4733 | -34.0787 | -6.2114
3.4794 | -34.0788 | -6.2221
3.4855 | -34.079 | -6.2324
3.4916 | -34.0791 | -6.2427
3.4977 | -34.0793 | -6.2538
3.5038 | -34.0794 | -6.2646
3.5099 | -34.0795 | -6.2748
3.5161 | -34.0797 | -6.285
3.5222 | -34.0798 | -6.2953
3.5283 | -34.0798 | -6.306
3.5344 | -34.08 | -6.3166
3.5405 | -34.0801 | -6.327
3.5466 | -34.0803 | -6.3379
3.5527 | -34.0804 | -6.3485
3.5588 | -34.0805 | -6.359
3.5649 | -34.0805 | -6.3696
3.571 | -34.0806 | -6.3801
3.5771 | -34.0807 | -6.3911
3.5832 | -34.0809 | -6.4008
3.5893 | -34.0811 | -6.4114
3.5954 | -34.0812 | -6.4221
3.6015 | -34.0814 | -6.4328
3.6076 | -34.0816 | -6.4433
3.6137 | -34.0817 | -6.4532
3.6198 | -34.082 | -6.4634
3.6259 | -34.0821 | -6.4742
3.632 | -34.0822 | -6.4846
3.6381 | -34.0824 | -6.4953
3.6442 | -34.0825 | -6.5058
3.6503 | -34.0827 | -6.5158
3.6565 | -34.0829 | -6.5264
3.6626 | -34.083 | -6.5369
3.6687 | -34.0832 | -6.5475
3.6748 | -34.0835 | -6.5579
3.6809 | -34.0837 | -6.5684
3.687 | -34.0839 | -6.5792
3.6931 | -34.0841 | -6.5893
3.6992 | -34.0843 | -6.5999
3.7053 | -34.0845 | -6.6102
3.7114 | -34.0846 | -6.6208
3.7175 | -34.0848 | -6.6314
3.7236 | -34.0851 | -6.6413
3.7297 | -34.0853 | -6.652
3.7358 | -34.0855 | -6.6628
3.7419 | -34.0856 | -6.6735
3.748 | -34.0859 | -6.684
3.7541 | -34.086 | -6.6944
3.7602 | -34.0863 | -6.7048
3.7663 | -34.0865 | -6.7151
3.7724 | -34.0867 | -6.7257
3.7785 | -34.087 | -6.736
3.7846 | -34.0872 | -6.7463
3.7907 | -34.0875 | -6.7571
3.7969 | -34.0877 | -6.7677
3.803 | -34.088 | -6.7778
3.8091 | -34.0882 | -6.7882
3.8152 | -34.0885 | -6.7986
3.8213 | -34.0887 | -6.8094
3.8274 | -34.0889 | -6.8195
3.8335 | -34.0893 | -6.8302
3.8396 | -34.0896 | -6.8407
3.8457 | -34.0899 | -6.8513
3.8518 | -34.0902 | -6.8617
3.8579 | -34.0905 | -6.8722
3.864 | -34.0908 | -6.8829
3.8701 | -34.0911 | -6.8938
3.8762 | -34.0913 | -6.9041
3.8823 | -34.0915 | -6.9148
3.8884 | -34.0919 | -6.9253
3.8945 | -34.0922 | -6.9361
3.9006 | -34.0925 | -6.947
3.9067 | -34.0928 | -6.958
3.9128 | -34.0931 | -6.9687
3.9189 | -34.0934 | -6.979
3.925 | -34.0936 | -6.9897
3.9311 | -34.0939 | -6.9997
3.9372 | -34.0942 | -7.0105
3.9434 | -34.0945 | -7.0211
3.9495 | -34.0948 | -7.0318
3.9556 | -34.0952 | -7.0422
3.9617 | -34.0954 | -7.0524
3.9678 | -34.0958 | -7.0626
3.9739 | -34.096 | -7.0732
3.98 | -34.0963 | -7.0838
3.9861 | -34.0966 | -7.0943
3.9922 | -34.0968 | -7.1049
3.9983 | -34.0971 | -7.115
4.0044 | -34.0974 | -7.1251
4.0105 | -34.0977 | -7.1357
4.0166 | -34.098 | -7.1463
4.0227 | -34.0983 | -7.1568
4.0288 | -34.0986 | -7.1678
4.0349 | -34.0988 | -7.1781
4.041 | -34.099 | -7.1885
4.0471 | -34.0993 | -7.1992
4.0532 | -34.0996 | -7.2093
4.0593 | -34.0999 | -7.2201
4.0654 | -34.1002 | -7.2303
4.0715 | -34.1004 | -7.2413
4.0776 | -34.1008 | -7.2517
4.0838 | -34.1011 | -7.2627
4.0899 | -34.1013 | -7.2729
4.096 | -34.1016 | -7.2834
4.1021 | -34.1019 | -7.2936
4.1082 | -34.1021 | -7.3039
4.1143 | -34.1024 | -7.314
4.1204 | -34.1027 | -7.3242
4.1265 | -34.103 | -7.3344
4.1326 | -34.1032 | -7.3445
4.1387 | -34.1036 | -7.3548
4.1448 | -34.1038 | -7.3653
4.1509 | -34.104 | -7.3758
4.157 | -34.1043 | -7.3866
4.1631 | -34.1046 | -7.3973
4.1692 | -34.1048 | -7.4075
4.1753 | -34.1051 | -7.4179
4.1814 | -34.1053 | -7.4279
4.1875 | -34.1056 | -7.4386
4.1936 | -34.1059 | -7.4492
4.1997 | -34.1061 | -7.4592
4.2058 | -34.1063 | -7.4699
4.2119 | -34.1066 | -7.4801
4.218 | -34.1068 | -7.4902
4.2241 | -34.107 | -7.5004
4.2303 | -34.1073 | -7.5106
4.2364 | -34.1075 | -7.5209
4.2425 | -34.1077 | -7.5313
4.2486 | -34.1079 | -7.5418
4.2547 | -34.1081 | -7.5519
4.2608 | -34.1083 | -7.5623
4.2669 | -34.1085 | -7.5723
4.273 | -34.1087 | -7.5828
4.2791 | -34.109 | -7.5931
4.2852 | -34.1092 | -7.6036
4.2913 | -34.1093 | -7.6136
4.2974 | -34.1095 | -7.6242
4.3035 | -34.1098 | -7.6348
4.3096 | -34.11 | -7.6453
4.3157 | -34.1102 | -7.6554
4.3218 | -34.1104 | -7.6656
4.3279 | -34.1106 | -7.6759
4.334 | -34.1108 | -7.6861
4.3401 | -34.111 | -7.6959
4.3462 | -34.1113 | -7.7063
4.3523 | -34.1115 | -7.7161
4.3584 | -34.1117 | -7.7259
4.3645 | -34.1119 | -7.7361
4.3707 | -34.1121 | -7.746
4.3768 | -34.1123 | -7.7562
4.3829 | -34.1125 | -7.7668
4.389 | -34.1127 | -7.777
4.3951 | -34.1129 | -7.787
4.4012 | -34.1131 | -7.7975
4.4073 | -34.1133 | -7.8076
4.4134 | -34.1135 | -7.8176
4.4195 | -34.1137 | -7.8278
4.4256 | -34.114 | -7.8378
4.4317 | -34.1142 | -7.8479
4.4378 | -34.1144 | -7.8577
4.4439 | -34.1146 | -7.8675
4.45 | -34.1147 | -7.8777
4.4561 | -34.1148 | -7.8878
4.4622 | -34.115 | -7.898
4.4683 | -34.1152 | -7.9081
4.4744 | -34.1154 | -7.9186
4.4805 | -34.1156 | -7.9283
4.4866 | -34.1158 | -7.9384
4.4927 | -34.116 | -7.9488
4.4988 | -34.1162 | -7.9588
4.5049 | -34.1164 | -7.9685
4.511 | -34.1167 | -7.9783
4.5172 | -34.1169 | -7.9884
4.5233 | -34.117 | -7.9989
4.5294 | -34.1173 | -8.0088
4.5355 | -34.1175 | -8.0189
4.5416 | -34.1176 | -8.029
4.5477 | -34.1179 | -8.0391
4.5538 | -34.118 | -8.0487
4.5599 | -34.1183 | -8.0587
4.566 | -34.1185 | -8.0688
4.5721 | -34.1187 | -8.0789
4.5782 | -34.1189 | -8.0892
4.5843 | -34.1191 | -8.0994
4.5904 | -34.1193 | -8.1095
4.5965 | -34.1196 | -8.1194
4.6026 | -34.1198 | -8.13
4.6087 | -34.1199 | -8.1399
4.6148 | -34.1201 | -8.1504
4.6209 | -34.1204 | -8.1604
4.627 | -34.1206 | -8.1707
4.6331 | -34.1207 | -8.1806
4.6392 | -34.1209 | -8.1908
4.6453 | -34.1212 | -8.2009
4.6514 | -34.1214 | -8.211
4.6576 | -34.1216 | -8.2215
4.6637 | -34.1217 | -8.2312
4.6698 | -34.1219 | -8.2411
4.6759 | -34.1221 | -8.2516
4.682 | -34.1223 | -8.2616
4.6881 | -34.1224 | -8.2716
4.6942 | -34.1227 | -8.2817
4.7003 | -34.1229 | -8.2919
4.7064 | -34.1231 | -8.3017
4.7125 | -34.1233 | -8.3121
4.7186 | -34.1235 | -8.3219
4.7247 | -34.1237 | -8.3321
4.7308 | -34.124 | -8.3422
4.7369 | -34.1242 | -8.3524
4.743 | -34.1244 | -8.3628
4.7491 | -34.1246 | -8.3733
4.7552 | -34.1248 | -8.3835
4.7613 | -34.1249 | -8.3938
4.7674 | -34.1251 | -8.4037
4.7735 | -34.1253 | -8.4143
4.7796 | -34.1256 | -8.4247
4.7857 | -34.1258 | -8.435
4.7918 | -34.126 | -8.445
4.7979 | -34.1263 | -8.455
4.8041 | -34.1265 | -8.4652
4.8102 | -34.1267 | -8.475
4.8163 | -34.1269 | -8.4851
4.8224 | -34.1271 | -8.4953
4.8285 | -34.1273 | -8.5051
4.8346 | -34.1276 | -8.515
4.8407 | -34.1278 | -8.5253
4.8468 | -34.128 | -8.5357
4.8529 | -34.1282 | -8.5462
4.859 | -34.1285 | -8.5566
4.8651 | -34.1287 | -8.5665
4.8712 | -34.1289 | -8.5765
4.8773 | -34.1291 | -8.5867
4.8834 | -34.1293 | -8.5969
4.8895 | -34.1296 | -8.6068
4.8956 | -34.1297 | -8.617
4.9017 | -34.1299 | -8.6269
4.9078 | -34.1301 | -8.6374
4.9139 | -34.1304 | -8.6472
4.92 | -34.1306 | -8.657
4.9261 | -34.1309 | -8.6668
4.9322 | -34.1311 | -8.677
4.9383 | -34.1314 | -8.6871
4.9445 | -34.1316 | -8.6973
4.9506 | -34.1318 | -8.7072
4.9567 | -34.1321 | -8.7172
4.9628 | -34.1323 | -8.7272
4.9689 | -34.1325 | -8.7378
4.975 | -34.1327 | -8.7482
4.9811 | -34.133 | -8.7581
4.9872 | -34.1332 | -8.7681
4.9933 | -34.1335 | -8.7785
4.9994 | -34.1338 | -8.7891
5.0055 | -34.134 | -8.7993
5.0116 | -34.1343 | -8.8099
5.0177 | -34.1346 | -8.8203
5.0238 | -34.1348 | -8.8308
5.0299 | -34.1352 | -8.8414
5.036 | -34.1354 | -8.8516
5.0421 | -34.1356 | -8.8617
5.0482 | -34.1359 | -8.8719
5.0543 | -34.1361 | -8.8818
5.0604 | -34.1364 | -8.892
5.0665 | -34.1367 | -8.9019
5.0726 | -34.137 | -8.9124
5.0787 | -34.1372 | -8.9229
5.0848 | -34.1375 | -8.9331
5.091 | -34.1377 | -8.9431
5.0971 | -34.138 | -8.953
5.1032 | -34.1383 | -8.9636
5.1093 | -34.1386 | -8.9738
5.1154 | -34.1388 | -8.9837
5.1215 | -34.139 | -8.9937
5.1276 | -34.1393 | -9.0037
5.1337 | -34.1396 | -9.0136
5.1398 | -34.1398 | -9.0235
5.1459 | -34.1401 | -9.034
5.152 | -34.1404 | -9.0443
5.1581 | -34.1406 | -9.0547
5.1642 | -34.1409 | -9.065
5.1703 | -34.1412 | -9.0752
5.1764 | -34.1415 | -9.0853
5.1825 | -34.1417 | -9.0952
5.1886 | -34.142 | -9.1051
5.1947 | -34.1423 | -9.1153
5.2008 | -34.1426 | -9.1256
5.2069 | -34.1428 | -9.1361
5.213 | -34.1431 | -9.1461
5.2191 | -34.1434 | -9.1563
5.2252 | -34.1437 | -9.1666
5.2314 | -34.1439 | -9.177
5.2375 | -34.1442 | -9.1869
5.2436 | -34.1444 | -9.1972
5.2497 | -34.1447 | -9.2071
5.2558 | -34.145 | -9.2177
5.2619 | -34.1453 | -9.2282
5.268 | -34.1456 | -9.2386
5.2741 | -34.1459 | -9.2489
5.2802 | -34.1462 | -9.2595
5.2863 | -34.1465 | -9.2697
5.2924 | -34.1468 | -9.2801
5.2985 | -34.1471 | -9.2905
5.3046 | -34.1474 | -9.3008
5.3107 | -34.1476 | -9.3113
5.3168 | -34.1479 | -9.3216
5.3229 | -34.1482 | -9.3319
5.329 | -34.1485 | -9.3421
5.3351 | -34.1488 | -9.3521
5.3412 | -34.149 | -9.3623
5.3473 | -34.1493 | -9.3727
5.3534 | -34.1497 | -9.3831
5.3595 | -34.1499 | -9.3934
5.3656 | -34.1503 | -9.4037
5.3717 | -34.1505 | -9.4143
5.3779 | -34.1508 | -9.425
5.384 | -34.1511 | -9.4352
5.3901 | -34.1514 | -9.4458
5.3962 | -34.1517 | -9.4559
5.4023 | -34.152 | -9.4664
5.4084 | -34.1523 | -9.4767
5.4145 | -34.1526 | -9.4867
5.4206 | -34.1529 | -9.4971
5.4267 | -34.1532 | -9.5077
5.4328 | -34.1535 | -9.5181
5.4389 | -34.1538 | -9.5281
5.445 | -34.154 | -9.5388
5.4511 | -34.1544 | -9.5489
5.4572 | -34.1547 | -9.5593
5.4633 | -34.1549 | -9.5696
5.4694 | -34.1552 | -9.58
5.4755 | -34.1555 | -9.5901
5.4816 | -34.1558 | -9.6004
5.4877 | -34.1561 | -9.6106
5.4938 | -34.1564 | -9.6207
5.4999 | -34.1567 | -9.6309
5.506 | -34.157 | -9.6412
5.5121 | -34.1573 | -9.6514
5.5183 | -34.1576 | -9.6615
5.5244 | -34.1579 | -9.6721
5.5305 | -34.1582 | -9.6823
5.5366 | -34.1585 | -9.6924
5.5427 | -34.1587 | -9.7028
5.5488 | -34.159 | -9.7132
5.5549 | -34.1593 | -9.7236
5.561 | -34.1596 | -9.7341
5.5671 | -34.1599 | -9.7443
5.5732 | -34.1602 | -9.7549
5.5793 | -34.1605 | -9.7656
5.5854 | -34.1609 | -9.7758
5.5915 | -34.1611 | -9.7865
5.5976 | -34.1614 | -9.7969
5.6037 | -34.1617 | -9.8073
5.6098 | -34.162 | -9.8176
5.6159 | -34.1623 | -9.8277
5.622 | -34.1626 | -9.8382
5.6281 | -34.1629 | -9.8486
5.6342 | -34.1632 | -9.8593
5.6403 | -34.1635 | -9.8697
5.6464 | -34.1638 | -9.8802
5.6525 | -34.1641 | -9.8906
5.6586 | -34.1644 | -9.9007
5.6648 | -34.1647 | -9.9111
5.6709 | -34.165 | -9.9216
5.677 | -34.1653 | -9.932
5.6831 | -34.1656 | -9.9424
5.6892 | -34.1659 | -9.953
5.6953 | -34.1662 | -9.9633
5.7014 | -34.1665 | -9.9738
5.7075 | -34.1668 | -9.9842
5.7136 | -34.1671 | -9.9945
5.7197 | -34.1673 | -10.0048
5.7258 | -34.1676 | -10.015
5.7319 | -34.1679 | -10.0254
5.738 | -34.1683 | -10.0358
5.7441 | -34.1686 | -10.0461
5.7502 | -34.1689 | -10.0562
5.7563 | -34.1692 | -10.0665
5.7624 | -34.1695 | -10.0771
5.7685 | -34.1698 | -10.0875
5.7746 | -34.1701 | -10.0982
5.7807 | -34.1703 | -10.1086
5.7868 | -34.1706 | -10.1192
5.7929 | -34.1709 | -10.1294
5.799 | -34.1712 | -10.1397
5.8052 | -34.1715 | -10.1499
5.8113 | -34.1718 | -10.1602
5.8174 | -34.1721 | -10.1704
5.8235 | -34.1724 | -10.1807
5.8296 | -34.1726 | -10.1912
5.8357 | -34.1729 | -10.2015
5.8418 | -34.1732 | -10.212
5.8479 | -34.1735 | -10.2225
5.854 | -34.1738 | -10.233
5.8601 | -34.1741 | -10.2432
5.8662 | -34.1744 | -10.2537
5.8723 | -34.1746 | -10.2639
5.8784 | -34.1749 | -10.2741
5.8845 | -34.1752 | -10.2844
5.8906 | -34.1755 | -10.2948
5.8967 | -34.1758 | -10.3049
5.9028 | -34.1761 | -10.3151
5.9089 | -34.1765 | -10.3253
5.915 | -34.1767 | -10.3354
5.9211 | -34.177 | -10.3457
5.9272 | -34.1773 | -10.3558
5.9333 | -34.1776 | -10.366
5.9394 | -34.1779 | -10.3763
5.9455 | -34.1782 | -10.3864
5.9517 | -34.1786 | -10.3965
5.9578 | -34.1789 | -10.407
5.9639 | -34.1792 | -10.4169
5.97 | -34.1795 | -10.4273
5.9761 | -34.1798 | -10.4375
5.9822 | -34.1801 | -10.4478
5.9883 | -34.1804 | -10.4581
5.9944 | -34.1807 | -10.4685
6.0005 | -34.181 | -10.479
6.0066 | -34.1813 | -10.489
6.0127 | -34.1817 | -10.499
6.0188 | -34.1819 | -10.5096
6.0249 | -34.1822 | -10.52
6.031 | -34.1825 | -10.5301
6.0371 | -34.1828 | -10.5403
6.0432 | -34.1831 | -10.5505
6.0493 | -34.1834 | -10.5606
6.0554 | -34.1838 | -10.5707
6.0615 | -34.1841 | -10.5809
6.0676 | -34.1844 | -10.5914
6.0737 | -34.1847 | -10.6018
6.0798 | -34.185 | -10.612
6.0859 | -34.1853 | -10.6221
6.0921 | -34.1856 | -10.6321
6.0982 | -34.1859 | -10.6424
6.1043 | -34.1863 | -10.6525
6.1104 | -34.1866 | -10.663
6.1165 | -34.1869 | -10.6727
6.1226 | -34.1872 | -10.6829
6.1287 | -34.1875 | -10.693
6.1348 | -34.1878 | -10.7032
6.1409 | -34.1881 | -10.7134
6.147 | -34.1884 | -10.7235
6.1531 | -34.1887 | -10.7336
6.1592 | -34.189 | -10.7437
6.1653 | -34.1893 | -10.7538
6.1714 | -34.1896 | -10.764
6.1775 | -34.1899 | -10.7743
6.1836 | -34.1902 | -10.7845
6.1897 | -34.1906 | -10.7946
6.1958 | -34.1908 | -10.8051
6.2019 | -34.1912 | -10.8154
6.208 | -34.1915 | -10.8256
6.2141 | -34.1918 | -10.8357
6.2202 | -34.1922 | -10.846
6.2263 | -34.1925 | -10.8561
6.2325 | -34.1927 | -10.8659
6.2386 | -34.1931 | -10.8757
6.2447 | -34.1934 | -10.8862
6.2508 | -34.1937 | -10.8963
6.2569 | -34.1941 | -10.9064
6.263 | -34.1943 | -10.9165
6.2691 | -34.1946 | -10.9266
6.2752 | -34.1949 | -10.9368
6.2813 | -34.1952 | -10.9469
6.2874 | -34.1955 | -10.9569
6.2935 | -34.1958 | -10.967
6.2996 | -34.1961 | -10.9771
6.3057 | -34.1964 | -10.9873
6.3118 | -34.1967 | -10.9975
6.3179 | -34.197 | -11.0076
6.324 | -34.1973 | -11.0176
6.3301 | -34.1976 | -11.0277
6.3362 | -34.1979 | -11.0377
6.3423 | -34.1983 | -11.048
6.3484 | -34.1986 | -11.0582
6.3545 | -34.1989 | -11.0679
6.3606 | -34.1991 | -11.0778
6.3667 | -34.1994 | -11.0877
6.3728 | -34.1997 | -11.0978
6.379 | -34.2 | -11.1079
6.3851 | -34.2004 | -11.1178
6.3912 | -34.2006 | -11.1278
6.3973 | -34.2009 | -11.1379
6.4034 | -34.2012 | -11.1477
6.4095 | -34.2015 | -11.1578
6.4156 | -34.2018 | -11.1679
6.4217 | -34.2021 | -11.1779
6.4278 | -34.2024 | -11.1881
6.4339 | -34.2027 | -11.198
6.44 | -34.2031 | -11.208
6.4461 | -34.2033 | -11.2183
6.4522 | -34.2037 | -11.2281
6.4583 | -34.2039 | -11.2383
6.4644 | -34.2042 | -11.2485
6.4705 | -34.2045 | -11.2581
6.4766 | -34.2048 | -11.2682
6.4827 | -34.2051 | -11.2783
6.4888 | -34.2054 | -11.2881
6.4949 | -34.2057 | -11.2979
6.501 | -34.206 | -11.308
6.5071 | -34.2063 | -11.3178
6.5132 | -34.2066 | -11.3278
6.5194 | -34.2069 | -11.3379
6.5255 | -34.2072 | -11.3476
6.5316 | -34.2075 | -11.3578
6.5377 | -34.2078 | -11.3677
6.5438 | -34.208 | -11.3775
6.5499 | -34.2084 | -11.3874
6.556 | -34.2087 | -11.3973
6.5621 | -34.2089 | -11.4076
6.5682 | -34.2092 | -11.4176
6.5743 | -34.2096 | -11.4275
6.5804 | -34.2099 | -11.4376
6.5865 | -34.2102 | -11.4476
6.5926 | -34.2105 | -11.4575
6.5987 | -34.2108 | -11.4672
6.6048 | -34.2111 | -11.4772
6.6109 | -34.2114 | -11.4871
6.617 | -34.2117 | -11.4972
6.6231 | -34.212 | -11.5071
6.6292 | -34.2123 | -11.5171
6.6353 | -34.2126 | -11.5273
6.6414 | -34.2129 | -11.5372
6.6475 | -34.2132 | -11.5471
6.6536 | -34.2135 | -11.5569
6.6597 | -34.2138 | -11.5666
6.6659 | -34.2141 | -11.5766
6.672 | -34.2144 | -11.5864
6.6781 | -34.2147 | -11.5963
6.6842 | -34.215 | -11.6061
6.6903 | -34.2154 | -11.616
6.6964 | -34.2157 | -11.626
6.7025 | -34.216 | -11.6357
6.7086 | -34.2162 | -11.6456
6.7147 | -34.2165 | -11.6556
6.7208 | -34.2168 | -11.6656
6.7269 | -34.2171 | -11.6755
6.733 | -34.2174 | -11.6852
6.7391 | -34.2177 | -11.6951
6.7452 | -34.218 | -11.7047
6.7513 | -34.2183 | -11.7144
6.7574 | -34.2186 | -11.7243
6.7635 | -34.2189 | -11.734
6.7696 | -34.2193 | -11.7438
6.7757 | -34.2196 | -11.7537
6.7818 | -34.2199 | -11.7636
6.7879 | -34.2202 | -11.7733
6.794 | -34.2205 | -11.7829
6.8001 | -34.2208 | -11.7927
6.8063 | -34.2211 | -11.8026
6.8124 | -34.2214 | -11.8126
6.8185 | -34.2217 | -11.8226
6.8246 | -34.222 | -11.8322
6.8307 | -34.2223 | -11.842
6.8368 | -34.2226 | -11.8521
6.8429 | -34.2229 | -11.8619
6.849 | -34.2232 | -11.8718
6.8551 | -34.2236 | -11.8816
6.8612 | -34.2239 | -11.8914
6.8673 | -34.2241 | -11.9011
6.8734 | -34.2244 | -11.9111
6.8795 | -34.2247 | -11.921
6.8856 | -34.2251 | -11.9308
6.8917 | -34.2254 | -11.9405
6.8978 | -34.2256 | -11.9501
6.9039 | -34.226 | -11.9596
6.91 | -34.2263 | -11.9695
6.9161 | -34.2266 | -11.9791
6.9222 | -34.2269 | -11.9889
6.9283 | -34.2273 | -11.9984
6.9344 | -34.2276 | -12.0082
6.9405 | -34.2279 | -12.018
6.9466 | -34.2282 | -12.028
6.9528 | -34.2286 | -12.038
6.9589 | -34.2289 | -12.0479
6.965 | -34.2292 | -12.0578
6.9711 | -34.2295 | -12.0677
6.9772 | -34.2299 | -12.0775
6.9833 | -34.2302 | -12.0872
6.9894 | -34.2305 | -12.0973
6.9955 | -34.2308 | -12.1068
7.0016 | -34.2311 | -12.1165
7.0077 | -34.2315 | -12.1263
7.0138 | -34.2318 | -12.136
7.0199 | -34.2321 | -12.1457
7.026 | -34.2324 | -12.1554
7.0321 | -34.2327 | -12.165
7.0382 | -34.2331 | -12.1744
7.0443 | -34.2334 | -12.1844
7.0504 | -34.2337 | -12.1942
7.0565 | -34.234 | -12.204
7.0626 | -34.2344 | -12.2135
7.0687 | -34.2347 | -12.2228
7.0748 | -34.235 | -12.2323
7.0809 | -34.2353 | -12.242
7.087 | -34.2357 | -12.2518
7.0932 | -34.2359 | -12.2616
7.0993 | -34.2363 | -12.2713
7.1054 | -34.2366 | -12.2809
7.1115 | -34.2369 | -12.2904
7.1176 | -34.2373 | -12.3
7.1237 | -34.2375 | -12.3097
7.1298 | -34.2379 | -12.3193
7.1359 | -34.2382 | -12.329
7.142 | -34.2385 | -12.3384
7.1481 | -34.2388 | -12.348
7.1542 | -34.2391 | -12.3578
7.1603 | -34.2395 | -12.3674
7.1664 | -34.2398 | -12.3771
7.1725 | -34.2402 | -12.3868
7.1786 | -34.2405 | -12.3965
7.1847 | -34.2408 | -12.4062
7.1908 | -34.2411 | -12.4161
7.1969 | -34.2414 | -12.4259
7.203 | -34.2418 | -12.4354
7.2091 | -34.2421 | -12.445
7.2152 | -34.2424 | -12.4547
7.2213 | -34.2428 | -12.4642
7.2274 | -34.2431 | -12.4738
7.2335 | -34.2434 | -12.4834
7.2397 | -34.2438 | -12.4931
7.2458 | -34.2441 | -12.5025
7.2519 | -34.2445 | -12.5122
7.258 | -34.2448 | -12.522
7.2641 | -34.2451 | -12.5316
7.2702 | -34.2454 | -12.5415
7.2763 | -34.2458 | -12.5512
7.2824 | -34.2461 | -12.561
7.2885 | -34.2464 | -12.5707
7.2946 | -34.2468 | -12.5803
7.3007 | -34.2471 | -12.5901
7.3068 | -34.2474 | -12.6
7.3129 | -34.2477 | -12.6099
7.319 | -34.2481 | -12.6196
7.3251 | -34.2484 | -12.629
7.3312 | -34.2487 | -12.6387
7.3373 | -34.2491 | -12.6481
7.3434 | -34.2494 | -12.6577
7.3495 | -34.2498 | -12.6672
7.3556 | -34.2501 | -12.677
7.3617 | -34.2504 | -12.6868
7.3678 | -34.2508 | -12.6964
7.3739 | -34.2511 | -12.7063
7.3801 | -34.2514 | -12.7162
7.3862 | -34.2518 | -12.7258
7.3923 | -34.2521 | -12.7356
7.3984 | -34.2524 | -12.7451
7.4045 | -34.2528 | -12.7547
7.4106 | -34.2532 | -12.7643
7.4167 | -34.2535 | -12.774
7.4228 | -34.2538 | -12.7838
7.4289 | -34.2542 | -12.7934
7.435 | -34.2545 | -12.8029
7.4411 | -34.2549 | -12.8127
7.4472 | -34.2552 | -12.8224
7.4533 | -34.2555 | -12.8319
7.4594 | -34.2559 | -12.8418
7.4655 | -34.2562 | -12.8512
7.4716 | -34.2566 | -12.861
7.4777 | -34.2569 | -12.8707
7.4838 | -34.2572 | -12.8802
7.4899 | -34.2576 | -12.8898
7.496 | -34.2579 | -12.8993
7.5021 | -34.2582 | -12.9089
7.5082 | -34.2586 | -12.9186
7.5143 | -34.2589 | -12.9283
7.5204 | -34.2593 | -12.9381
7.5266 | -34.2596 | -12.9476
7.5327 | -34.2599 | -12.9571
7.5388 | -34.2602 | -12.9671
7.5449 | -34.2606 | -12.977
7.551 | -34.261 | -12.9865
7.5571 | -34.2613 | -12.9962
7.5632 | -34.2617 | -13.006
7.5693 | -34.262 | -13.0156
7.5754 | -34.2624 | -13.0251
7.5815 | -34.2627 | -13.0348
7.5876 | -34.263 | -13.0447
7.5937 | -34.2634 | -13.0545
7.5998 | -34.2638 | -13.0641
7.6059 | -34.2642 | -13.0739
7.612 | -34.2645 | -13.0836
7.6181 | -34.2648 | -13.0932
7.6242 | -34.2652 | -13.1031
7.6303 | -34.2655 | -13.1131
7.6364 | -34.2658 | -13.1228
7.6425 | -34.2662 | -13.1325
7.6486 | -34.2665 | -13.1422
7.6547 | -34.2669 | -13.1518
7.6608 | -34.2672 | -13.1615
7.667 | -34.2675 | -13.1713
7.6731 | -34.2679 | -13.181
7.6792 | -34.2683 | -13.191
7.6853 | -34.2686 | -13.2008
7.6914 | -34.269 | -13.2106
7.6975 | -34.2693 | -13.2204
7.7036 | -34.2697 | -13.2301
7.7097 | -34.27 | -13.24
7.7158 | -34.2704 | -13.2498
7.7219 | -34.2707 | -13.2594
7.728 | -34.2711 | -13.2692
7.7341 | -34.2714 | -13.2791
7.7402 | -34.2718 | -13.2889
7.7463 | -34.2721 | -13.2986
7.7524 | -34.2725 | -13.3086
7.7585 | -34.2728 | -13.3184
7.7646 | -34.2732 | -13.3281
7.7707 | -34.2735 | -13.3381
7.7768 | -34.2738 | -13.348
7.7829 | -34.2742 | -13.358
7.789 | -34.2745 | -13.368
7.7951 | -34.2749 | -13.3778
7.8012 | -34.2752 | -13.3875
7.8073 | -34.2756 | -13.3976
7.8135 | -34.276 | -13.4073
7.8196 | -34.2763 | -13.4173
7.8257 | -34.2767 | -13.4273
7.8318 | -34.277 | -13.4372
7.8379 | -34.2774 | -13.4469
7.844 | -34.2777 | -13.4566
7.8501 | -34.2781 | -13.4665
7.8562 | -34.2784 | -13.4765
7.8623 | -34.2788 | -13.4862
7.8684 | -34.2791 | -13.4961
7.8745 | -34.2795 | -13.506
7.8806 | -34.2798 | -13.516
7.8867 | -34.2802 | -13.5255
7.8928 | -34.2805 | -13.5351
7.8989 | -34.2809 | -13.5449
7.905 | -34.2812 | -13.5548
7.9111 | -34.2816 | -13.5647
7.9172 | -34.2819 | -13.5745
7.9233 | -34.2823 | -13.5844
7.9294 | -34.2826 | -13.5943
7.9355 | -34.283 | -13.6039
7.9416 | -34.2833 | -13.6138
7.9477 | -34.2837 | -13.6235
7.9539 | -34.284 | -13.6332
7.96 | -34.2844 | -13.6432
7.9661 | -34.2847 | -13.653
7.9722 | -34.2851 | -13.6627
7.9783 | -34.2855 | -13.6728
7.9844 | -34.2858 | -13.6826
7.9905 | -34.2862 | -13.6926
7.9966 | -34.2865 | -13.7026
8.0027 | -34.2869 | -13.7124
8.0088 | -34.2872 | -13.7222
8.0149 | -34.2875 | -13.7321
8.021 | -34.2879 | -13.742
8.0271 | -34.2883 | -13.7521
8.0332 | -34.2886 | -13.7619
8.0393 | -34.289 | -13.7719
8.0454 | -34.2893 | -13.782
8.0515 | -34.2897 | -13.7919
8.0576 | -34.2901 | -13.802
8.0637 | -34.2904 | -13.8117
8.0698 | -34.2907 | -13.8214
8.0759 | -34.2911 | -13.8312
8.082 | -34.2914 | -13.8409
8.0881 | -34.2918 | -13.8509
8.0942 | -34.2922 | -13.8608
8.1004 | -34.2925 | -13.8707
8.1065 | -34.2929 | -13.8805
8.1126 | -34.2932 | -13.8903
8.1187 | -34.2935 | -13.9003
8.1248 | -34.2939 | -13.9099
8.1309 | -34.2943 | -13.9196
8.137 | -34.2946 | -13.9293
8.1431 | -34.295 | -13.9394
8.1492 | -34.2953 | -13.9492
8.1553 | -34.2957 | -13.9591
8.1614 | -34.296 | -13.9691
8.1675 | -34.2963 | -13.979
8.1736 | -34.2967 | -13.989
8.1797 | -34.297 | -13.9989
8.1858 | -34.2974 | -14.0089
8.1919 | -34.2977 | -14.0187
8.198 | -34.2981 | -14.0284
8.2041 | -34.2985 | -14.0383
8.2102 | -34.2988 | -14.0482
8.2163 | -34.2992 | -14.0579
8.2224 | -34.2995 | -14.0677
8.2285 | -34.2999 | -14.0773
8.2346 | -34.3002 | -14.0871
8.2408 | -34.3006 | -14.0968
8.2469 | -34.301 | -14.1067
8.253 | -34.3013 | -14.1164
8.2591 | -34.3017 | -14.1263
8.2652 | -34.3021 | -14.1362
8.2713 | -34.3024 | -14.1461
8.2774 | -34.3028 | -14.1561
8.2835 | -34.3031 | -14.1662
8.2896 | -34.3035 | -14.1759
8.2957 | -34.3039 | -14.1858
8.3018 | -34.3042 | -14.1957
8.3079 | -34.3046 | -14.2053
8.314 | -34.305 | -14.2152
8.3201 | -34.3053 | -14.225
8.3262 | -34.3057 | -14.2349
8.3323 | -34.306 | -14.2447
8.3384 | -34.3064 | -14.2547
8.3445 | -34.3067 | -14.2645
8.3506 | -34.3071 | -14.2744
8.3567 | -34.3075 | -14.2843
8.3628 | -34.3078 | -14.2942
8.3689 | -34.3082 | -14.3038
8.375 | -34.3085 | -14.3138
8.3812 | -34.3089 | -14.3235
8.3873 | -34.3092 | -14.3334
8.3934 | -34.3096 | -14.3433
8.3995 | -34.3099 | -14.353
8.4056 | -34.3103 | -14.3629
8.4117 | -34.3106 | -14.3728
8.4178 | -34.311 | -14.3827
8.4239 | -34.3113 | -14.3923
8.43 | -34.3117 | -14.4022
8.4361 | -34.312 | -14.4121
8.4422 | -34.3124 | -14.4218
8.4483 | -34.3127 | -14.4316
8.4544 | -34.3131 | -14.4414
8.4605 | -34.3135 | -14.4511
8.4666 | -34.3138 | -14.4612
8.4727 | -34.3141 | -14.4711
8.4788 | -34.3145 | -14.4809
8.4849 | -34.3149 | -14.4907
8.491 | -34.3152 | -14.5005
8.4971 | -34.3156 | -14.5101
8.5032 | -34.3159 | -14.5198
8.5093 | -34.3163 | -14.5294
8.5154 | -34.3166 | -14.5393
8.5215 | -34.317 | -14.549
8.5277 | -34.3173 | -14.5589
8.5338 | -34.3177 | -14.5686
8.5399 | -34.318 | -14.5784
8.546 | -34.3184 | -14.588
8.5521 | -34.3188 | -14.5978
8.5582 | -34.3191 | -14.6075
8.5643 | -34.3195 | -14.617
8.5704 | -34.3198 | -14.6266
8.5765 | -34.3202 | -14.6365
8.5826 | -34.3206 | -14.646
8.5887 | -34.3209 | -14.6556
8.5948 | -34.3213 | -14.6655
8.6009 | -34.3217 | -14.675
8.607 | -34.322 | -14.6849
8.6131 | -34.3224 | -14.6948
8.6192 | -34.3228 | -14.7043
8.6253 | -34.3231 | -14.714
8.6314 | -34.3235 | -14.7237
8.6375 | -34.3238 | -14.7335
8.6436 | -34.3242 | -14.7432
8.6497 | -34.3245 | -14.7528
8.6558 | -34.3248 | -14.7625
8.6619 | -34.3252 | -14.7723
8.6681 | -34.3255 | -14.7822
8.6742 | -34.3259 | -14.7918
8.6803 | -34.3263 | -14.8016
8.6864 | -34.3266 | -14.811
8.6925 | -34.327 | -14.8205
8.6986 | -34.3273 | -14.8303
8.7047 | -34.3277 | -14.8402
8.7108 | -34.328 | -14.8499
8.7169 | -34.3284 | -14.8596
8.723 | -34.3287 | -14.8691
8.7291 | -34.3291 | -14.8786
8.7352 | -34.3295 | -14.8885
8.7413 | -34.3298 | -14.8983
8.7474 | -34.3301 | -14.9078
8.7535 | -34.3305 | -14.9175
8.7596 | -34.3309 | -14.9272
8.7657 | -34.3313 | -14.937
8.7718 | -34.3316 | -14.9468
8.7779 | -34.332 | -14.9565
8.784 | -34.3323 | -14.9663
8.7901 | -34.3327 | -14.976
8.7962 | -34.3331 | -14.9854
8.8023 | -34.3334 | -14.9951
8.8084 | -34.3338 | -15.0047
8.8146 | -34.3341 | -15.0143
8.8207 | -34.3345 | -15.0239
8.8268 | -34.3349 | -15.0333
8.8329 | -34.3352 | -15.0429
8.839 | -34.3356 | -15.0526
8.8451 | -34.3359 | -15.0621
8.8512 | -34.3363 | -15.0717
8.8573 | -34.3366 | -15.0814
8.8634 | -34.337 | -15.0908
8.8695 | -34.3374 | -15.1007
8.8756 | -34.3377 | -15.1103
8.8817 | -34.3381 | -15.1198
8.8878 | -34.3384 | -15.1297
8.8939 | -34.3388 | -15.1393
8.9 | -34.3392 | -15.1489
8.9061 | -34.3395 | -15.1585
8.9122 | -34.3399 | -15.1681
8.9183 | -34.3403 | -15.1777
8.9244 | -34.3406 | -15.1875
8.9305 | -34.341 | -15.1969
8.9366 | -34.3414 | -15.2065
8.9427 | -34.3417 | -15.2161
8.9488 | -34.3421 | -15.2256
8.955 | -34.3425 | -15.2352
8.9611 | -34.3428 | -15.2449
8.9672 | -34.3433 | -15.2545
8.9733 | -34.3436 | -15.2639
8.9794 | -34.344 | -15.2733
8.9855 | -34.3444 | -15.283
8.9916 | -34.3447 | -15.2928
8.9977 | -34.3451 | -15.3024
9.0038 | -34.3455 | -15.3118
9.0099 | -34.3459 | -15.3215
9.016 | -34.3462 | -15.3312
9.0221 | -34.3466 | -15.3406
9.0282 | -34.347 | -15.35
9.0343 | -34.3474 | -15.3597
9.0404 | -34.3478 | -15.3692
9.0465 | -34.3482 | -15.3789
9.0526 | -34.3486 | -15.3883
9.0587 | -34.349 | -15.3979
9.0648 | -34.3493 | -15.4074
9.0709 | -34.3497 | -15.4169
9.077 | -34.3501 | -15.4265
9.0831 | -34.3505 | -15.4363
9.0892 | -34.3508 | -15.4459
9.0953 | -34.3512 | -15.4555
9.1015 | -34.3516 | -15.4651
9.1076 | -34.352 | -15.4745
9.1137 | -34.3524 | -15.484
9.1198 | -34.3528 | -15.4935
9.1259 | -34.3532 | -15.5031
9.132 | -34.3535 | -15.5129
9.1381 | -34.3539 | -15.5223
9.1442 | -34.3543 | -15.5316
9.1503 | -34.3547 | -15.5412
9.1564 | -34.3551 | -15.5507
9.1625 | -34.3555 | -15.5603
9.1686 | -34.3559 | -15.5698
9.1747 | -34.3563 | -15.5793
9.1808 | -34.3567 | -15.5888
9.1869 | -34.3571 | -15.5983
9.193 | -34.3574 | -15.6079
9.1991 | -34.3578 | -15.6175
9.2052 | -34.3582 | -15.6271
9.2113 | -34.3586 | -15.6367
9.2174 | -34.359 | -15.6461
9.2235 | -34.3594 | -15.6556
9.2296 | -34.3598 | -15.6652
9.2357 | -34.3602 | -15.6749
9.2419 | -34.3606 | -15.6844
9.248 | -34.3611 | -15.6938
9.2541 | -34.3615 | -15.7035
9.2602 | -34.3619 | -15.7131
9.2663 | -34.3623 | -15.7223
9.2724 | -34.3627 | -15.7319
9.2785 | -34.3631 | -15.7414
9.2846 | -34.3635 | -15.7507
9.2907 | -34.3639 | -15.7602
9.2968 | -34.3643 | -15.7699
9.3029 | -34.3646 | -15.7795
9.309 | -34.365 | -15.7891
9.3151 | -34.3654 | -15.7985
9.3212 | -34.3658 | -15.808
9.3273 | -34.3662 | -15.8175
9.3334 | -34.3666 | -15.8268
9.3395 | -34.367 | -15.8362
9.3456 | -34.3674 | -15.8455
9.3517 | -34.3677 | -15.8551
9.3578 | -34.3681 | -15.8646
9.3639 | -34.3686 | -15.874
9.37 | -34.369 | -15.8836
9.3761 | -34.3694 | -15.8933
9.3822 | -34.3698 | -15.9028
9.3884 | -34.3702 | -15.912
9.3945 | -34.3706 | -15.9217
9.4006 | -34.371 | -15.9312
9.4067 | -34.3714 | -15.9408
9.4128 | -34.3718 | -15.9503
9.4189 | -34.3722 | -15.9599
9.425 | -34.3726 | -15.969
9.4311 | -34.373 | -15.9785
9.4372 | -34.3735 | -15.988
9.4433 | -34.3739 | -15.9976
9.4494 | -34.3743 | -16.0068
9.4555 | -34.3747 | -16.0164
9.4616 | -34.3751 | -16.0258
9.4677 | -34.3755 | -16.0352
9.4738 | -34.3759 | -16.0447
9.4799 | -34.3764 | -16.0542
9.486 | -34.3768 | -16.0636
9.4921 | -34.3772 | -16.0731
9.4982 | -34.3776 | -16.0824
9.5043 | -34.378 | -16.0918
9.5104 | -34.3784 | -16.1011
9.5165 | -34.3788 | -16.1105
9.5226 | -34.3792 | -16.1201
9.5288 | -34.3796 | -16.1293
9.5349 | -34.38 | -16.1387
9.541 | -34.3805 | -16.148
9.5471 | -34.3809 | -16.1575
9.5532 | -34.3813 | -16.167
9.5593 | -34.3817 | -16.1765
9.5654 | -34.3821 | -16.1859
9.5715 | -34.3825 | -16.1953
9.5776 | -34.3829 | -16.2051
9.5837 | -34.3833 | -16.2145
9.5898 | -34.3836 | -16.224
9.5959 | -34.3841 | -16.2334
9.602 | -34.3845 | -16.2427
9.6081 | -34.3849 | -16.252
9.6142 | -34.3853 | -16.2615
9.6203 | -34.3857 | -16.2709
9.6264 | -34.3861 | -16.2802
9.6325 | -34.3865 | -16.2897
9.6386 | -34.3869 | -16.299
9.6447 | -34.3873 | -16.3084
9.6508 | -34.3877 | -16.3177
9.6569 | -34.3881 | -16.3271
9.663 | -34.3885 | -16.3366
9.6691 | -34.3889 | -16.3462
9.6753 | -34.3894 | -16.3556
9.6814 | -34.3898 | -16.3649
9.6875 | -34.3902 | -16.3741
9.6936 | -34.3906 | -16.3835
9.6997 | -34.391 | -16.3931
9.7058 | -34.3915 | -16.4023
9.7119 | -34.3919 | -16.4118
9.718 | -34.3923 | -16.4214
9.7241 | -34.3927 | -16.4305
9.7302 | -34.3931 | -16.44
9.7363 | -34.3935 | -16.4493
9.7424 | -34.394 | -16.4587
9.7485 | -34.3944 | -16.4682
9.7546 | -34.3948 | -16.4774
9.7607 | -34.3951 | -16.4871
9.7668 | -34.3956 | -16.4965
9.7729 | -34.396 | -16.5057
9.779 | -34.3964 | -16.5151
9.7851 | -34.3968 | -16.5248
9.7912 | -34.3973 | -16.5342
9.7973 | -34.3977 | -16.5437
9.8034 | -34.3981 | -16.5529
9.8095 | -34.3985 | -16.5624
9.8157 | -34.399 | -16.5716
9.8218 | -34.3994 | -16.5811
9.8279 | -34.3998 | -16.5903
9.834 | -34.4002 | -16.5998
9.8401 | -34.4006 | -16.6092
9.8462 | -34.401 | -16.6183
9.8523 | -34.4014 | -16.6276
9.8584 | -34.4018 | -16.637
9.8645 | -34.4022 | -16.6464
9.8706 | -34.4026 | -16.6559
9.8767 | -34.403 | -16.6653
9.8828 | -34.4034 | -16.6745
9.8889 | -34.4038 | -16.6838
9.895 | -34.4042 | -16.6931
9.9011 | -34.4046 | -16.7024
9.9072 | -34.4051 | -16.7118
9.9133 | -34.4054 | -16.7211
9.9194 | -34.4059 | -16.7304
9.9255 | -34.4063 | -16.7396
9.9316 | -34.4067 | -16.7488
9.9377 | -34.4071 | -16.7582
9.9438 | -34.4075 | -16.7676
9.9499 | -34.4079 | -16.7772
9.956 | -34.4083 | -16.7867
9.9622 | -34.4087 | -16.7961
9.9683 | -34.4092 | -16.8054
9.9744 | -34.4095 | -16.8149
9.9805 | -34.41 | -16.8242
9.9866 | -34.4104 | -16.8334
9.9927 | -34.4108 | -16.8427
9.9988 | -34.4112 | -16.8521
10.0049 | -34.4116 | -16.8614
10.011 | -34.412 | -16.8707
10.0171 | -34.4124 | -16.88
10.0232 | -34.4129 | -16.8892
10.0293 | -34.4133 | -16.8986
10.0354 | -34.4138 | -16.908
10.0415 | -34.4142 | -16.9173
10.0476 | -34.4146 | -16.9267
10.0537 | -34.415 | -16.9362
10.0598 | -34.4154 | -16.9456
10.0659 | -34.4158 | -16.9549
10.072 | -34.4162 | -16.964
10.0781 | -34.4167 | -16.9733
10.0842 | -34.4171 | -16.9826
10.0903 | -34.4175 | -16.9919
10.0964 | -34.4179 | -17.001
10.1026 | -34.4183 | -17.0103
10.1087 | -34.4187 | -17.0196
10.1148 | -34.4191 | -17.0291
10.1209 | -34.4195 | -17.0382
10.127 | -34.4199 | -17.0473
10.1331 | -34.4204 | -17.0565
10.1392 | -34.4208 | -17.0656
10.1453 | -34.4212 | -17.0749
10.1514 | -34.4216 | -17.0843
10.1575 | -34.422 | -17.0937
10.1636 | -34.4224 | -17.1027
10.1697 | -34.4228 | -17.1122
10.1758 | -34.4233 | -17.1213
10.1819 | -34.4237 | -17.1306
10.188 | -34.424 | -17.14
10.1941 | -34.4244 | -17.1494
10.2002 | -34.4248 | -17.1588
10.2063 | -34.4252 | -17.168
10.2124 | -34.4256 | -17.1774
10.2185 | -34.4261 | -17.1868
10.2246 | -34.4265 | -17.1959
10.2307 | -34.4269 | -17.2051
10.2368 | -34.4273 | -17.2141
10.2429 | -34.4277 | -17.2234
10.2491 | -34.4281 | -17.2328
10.2552 | -34.4285 | -17.242
10.2613 | -34.4289 | -17.2513
10.2674 | -34.4293 | -17.2606
10.2735 | -34.4298 | -17.2702
10.2796 | -34.4302 | -17.2795
10.2857 | -34.4306 | -17.2888
10.2918 | -34.431 | -17.2982
10.2979 | -34.4314 | -17.3075
10.304 | -34.4319 | -17.3166
10.3101 | -34.4322 | -17.3258
10.3162 | -34.4326 | -17.3351
10.3223 | -34.433 | -17.3441
10.3284 | -34.4334 | -17.3533
10.3345 | -34.4338 | -17.3625
10.3406 | -34.4342 | -17.3717
10.3467 | -34.4346 | -17.3811
10.3528 | -34.435 | -17.3902
10.3589 | -34.4354 | -17.3994
10.365 | -34.4358 | -17.4087
10.3711 | -34.4362 | -17.4182
10.3772 | -34.4366 | -17.4273
10.3833 | -34.437 | -17.4366
10.3895 | -34.4375 | -17.4458
10.3956 | -34.4379 | -17.4551
10.4017 | -34.4383 | -17.4641
10.4078 | -34.4387 | -17.4734
10.4139 | -34.4391 | -17.4827
10.42 | -34.4395 | -17.492
10.4261 | -34.4399 | -17.5012
10.4322 | -34.4403 | -17.5104
10.4383 | -34.4407 | -17.5195
10.4444 | -34.4411 | -17.5284
10.4505 | -34.4415 | -17.5377
10.4566 | -34.4419 | -17.5468
10.4627 | -34.4423 | -17.556
10.4688 | -34.4427 | -17.565
10.4749 | -34.4431 | -17.5743
10.481 | -34.4435 | -17.5839
10.4871 | -34.444 | -17.593
10.4932 | -34.4444 | -17.6021
10.4993 | -34.4448 | -17.6114
10.5054 | -34.4452 | -17.6207
10.5115 | -34.4456 | -17.63
10.5176 | -34.446 | -17.6392
10.5237 | -34.4464 | -17.6485
10.5298 | -34.4468 | -17.6576
10.536 | -34.4472 | -17.6669
10.5421 | -34.4476 | -17.6762
10.5482 | -34.448 | -17.6853
10.5543 | -34.4484 | -17.6945
10.5604 | -34.4488 | -17.7035
10.5665 | -34.4492 | -17.7125
10.5726 | -34.4496 | -17.7216
10.5787 | -34.45 | -17.7306
10.5848 | -34.4504 | -17.7395
10.5909 | -34.4509 | -17.7488
10.597 | -34.4512 | -17.758
10.6031 | -34.4516 | -17.7672
10.6092 | -34.452 | -17.7764
10.6153 | -34.4525 | -17.7857
10.6214 | -34.4529 | -17.7946
10.6275 | -34.4532 | -17.8038
10.6336 | -34.4536 | -17.8129
10.6397 | -34.454 | -17.8222
10.6458 | -34.4544 | -17.8312
10.6519 | -34.4548 | -17.8404
10.658 | -34.4552 | -17.8496
10.6641 | -34.4556 | -17.859
10.6702 | -34.456 | -17.8682
10.6764 | -34.4564 | -17.8775
10.6825 | -34.4568 | -17.8866
10.6886 | -34.4572 | -17.8956
10.6947 | -34.4576 | -17.9048
10.7008 | -34.4581 | -17.914
10.7069 | -34.4585 | -17.9233
10.713 | -34.4589 | -17.9325
10.7191 | -34.4593 | -17.9416
10.7252 | -34.4597 | -17.9504
10.7313 | -34.4601 | -17.9596
10.7374 | -34.4604 | -17.9688
10.7435 | -34.4608 | -17.978
10.7496 | -34.4612 | -17.987
10.7557 | -34.4616 | -17.9963
10.7618 | -34.462 | -18.0052
10.7679 | -34.4624 | -18.0143
10.774 | -34.4628 | -18.0235
10.7801 | -34.4632 | -18.0327
10.7862 | -34.4636 | -18.0416
10.7923 | -34.464 | -18.0509
10.7984 | -34.4644 | -18.0598
10.8045 | -34.4648 | -18.0688
10.8106 | -34.4652 | -18.078
10.8168 | -34.4656 | -18.0871
10.8229 | -34.466 | -18.0962
10.829 | -34.4664 | -18.1053
10.8351 | -34.4668 | -18.114
10.8412 | -34.4672 | -18.123
10.8473 | -34.4677 | -18.1321
10.8534 | -34.4681 | -18.1411
10.8595 | -34.4685 | -18.1502
10.8656 | -34.4689 | -18.1593
10.8717 | -34.4693 | -18.1685
10.8778 | -34.4697 | -18.1776
10.8839 | -34.4701 | -18.1868
10.89 | -34.4705 | -18.196
10.8961 | -34.4709 | -18.2051
10.9022 | -34.4713 | -18.2141
10.9083 | -34.4717 | -18.2231
10.9144 | -34.4721 | -18.2321
10.9205 | -34.4725 | -18.2412
10.9266 | -34.4729 | -18.2502
10.9327 | -34.4733 | -18.2591
10.9388 | -34.4737 | -18.2681
10.9449 | -34.4741 | -18.2772
10.951 | -34.4745 | -18.2862
10.9571 | -34.4749 | -18.2954
10.9633 | -34.4753 | -18.3045
10.9694 | -34.4757 | -18.3138
10.9755 | -34.4762 | -18.3229
10.9816 | -34.4765 | -18.332
10.9877 | -34.477 | -18.341
10.9938 | -34.4774 | -18.3502
10.9999 | -34.4778 | -18.3591
11.006 | -34.4782 | -18.3681
11.0121 | -34.4786 | -18.3771
11.0182 | -34.479 | -18.3862
11.0243 | -34.4794 | -18.3952
11.0304 | -34.4798 | -18.4043
11.0365 | -34.4802 | -18.4132
11.0426 | -34.4806 | -18.4224
11.0487 | -34.481 | -18.4316
11.0548 | -34.4815 | -18.4406
11.0609 | -34.4819 | -18.4498
11.067 | -34.4823 | -18.4589
11.0731 | -34.4827 | -18.4677
11.0792 | -34.4831 | -18.4768
11.0853 | -34.4835 | -18.4859
11.0914 | -34.484 | -18.495
11.0975 | -34.4843 | -18.504
11.1037 | -34.4847 | -18.5132
11.1098 | -34.4851 | -18.5225
11.1159 | -34.4856 | -18.5314
11.122 | -34.486 | -18.5403
11.1281 | -34.4864 | -18.5493
11.1342 | -34.4868 | -18.5583
11.1403 | -34.4872 | -18.5676
11.1464 | -34.4876 | -18.5767
11.1525 | -34.488 | -18.5853
11.1586 | -34.4884 | -18.5944
11.1647 | -34.4888 | -18.6034
11.1708 | -34.4893 | -18.6124
11.1769 | -34.4896 | -18.6215
11.183 | -34.4901 | -18.6303
11.1891 | -34.4905 | -18.6393
11.1952 | -34.4909 | -18.6483
11.2013 | -34.4913 | -18.6575
11.2074 | -34.4917 | -18.6666
11.2135 | -34.4921 | -18.6759
11.2196 | -34.4926 | -18.685
11.2257 | -34.493 | -18.6942
11.2318 | -34.4934 | -18.7031
11.2379 | -34.4938 | -18.7124
11.244 | -34.4943 | -18.7216
11.2502 | -34.4947 | -18.7308
11.2563 | -34.4951 | -18.7398
11.2624 | -34.4955 | -18.7488
11.2685 | -34.4959 | -18.7577
11.2746 | -34.4963 | -18.7666
11.2807 | -34.4968 | -18.7757
11.2868 | -34.4972 | -18.785
11.2929 | -34.4976 | -18.7936
11.299 | -34.498 | -18.8027
11.3051 | -34.4984 | -18.8119
11.3112 | -34.4989 | -18.8211
11.3173 | -34.4993 | -18.83
11.3234 | -34.4997 | -18.8389
11.3295 | -34.5001 | -18.8481
11.3356 | -34.5005 | -18.8572
11.3417 | -34.5009 | -18.866
11.3478 | -34.5013 | -18.8752
11.3539 | -34.5017 | -18.8844
11.36 | -34.5021 | -18.8931
11.3661 | -34.5026 | -18.902
11.3722 | -34.503 | -18.9109
11.3783 | -34.5034 | -18.92
11.3844 | -34.5038 | -18.9291
11.3906 | -34.5043 | -18.9381
11.3967 | -34.5047 | -18.947
11.4028 | -34.5051 | -18.9562
11.4089 | -34.5055 | -18.9651
11.415 | -34.5059 | -18.9741
11.4211 | -34.5063 | -18.9831
11.4272 | -34.5067 | -18.9923
11.4333 | -34.5071 | -19.0011
11.4394 | -34.5075 | -19.0101
11.4455 | -34.5079 | -19.019
11.4516 | -34.5083 | -19.0281
11.4577 | -34.5088 | -19.0371
11.4638 | -34.5092 | -19.046
11.4699 | -34.5096 | -19.055
11.476 | -34.51 | -19.0642
11.4821 | -34.5105 | -19.0732
11.4882 | -34.5109 | -19.0823
11.4943 | -34.5113 | -19.0915
11.5004 | -34.5117 | -19.1003
11.5065 | -34.5121 | -19.1092
11.5126 | -34.5125 | -19.1181
11.5187 | -34.513 | -19.1272
11.5248 | -34.5134 | -19.1363
11.5309 | -34.5138 | -19.1453
11.5371 | -34.5142 | -19.1542
11.5432 | -34.5147 | -19.1633
11.5493 | -34.5151 | -19.1726
11.5554 | -34.5155 | -19.1816
11.5615 | -34.5159 | -19.1907
11.5676 | -34.5164 | -19.1997
11.5737 | -34.5168 | -19.2088
11.5798 | -34.5173 | -19.2177
11.5859 | -34.5177 | -19.2266
11.592 | -34.5181 | -19.2357
11.5981 | -34.5185 | -19.2446
11.6042 | -34.519 | -19.2537
11.6103 | -34.5194 | -19.2626
11.6164 | -34.5198 | -19.2715
11.6225 | -34.5202 | -19.2804
11.6286 | -34.5206 | -19.2894
11.6347 | -34.521 | -19.2985
11.6408 | -34.5215 | -19.3074
11.6469 | -34.5219 | -19.3163
11.653 | -34.5223 | -19.325
11.6591 | -34.5227 | -19.334
11.6652 | -34.5231 | -19.3429
11.6713 | -34.5235 | -19.3517
11.6775 | -34.524 | -19.3607
11.6836 | -34.5244 | -19.3695
11.6897 | -34.5249 | -19.3786
11.6958 | -34.5253 | -19.3876
11.7019 | -34.5257 | -19.3967
11.708 | -34.5262 | -19.4056
11.7141 | -34.5266 | -19.4144
11.7202 | -34.5271 | -19.4235
11.7263 | -34.5274 | -19.4324
11.7324 | -34.5279 | -19.4414
11.7385 | -34.5283 | -19.4504
11.7446 | -34.5287 | -19.4593
11.7507 | -34.5292 | -19.468
11.7568 | -34.5296 | -19.4769
11.7629 | -34.53 | -19.4859
11.769 | -34.5304 | -19.4948
11.7751 | -34.5308 | -19.5038
11.7812 | -34.5314 | -19.5126
11.7873 | -34.5318 | -19.5216
11.7934 | -34.5322 | -19.5305
11.7995 | -34.5327 | -19.5394
11.8056 | -34.5331 | -19.5482
11.8117 | -34.5335 | -19.5572
11.8178 | -34.534 | -19.566
11.824 | -34.5344 | -19.5748
11.8301 | -34.5348 | -19.5836
11.8362 | -34.5352 | -19.5923
11.8423 | -34.5357 | -19.6012
11.8484 | -34.5361 | -19.6102
11.8545 | -34.5366 | -19.6191
11.8606 | -34.537 | -19.628
11.8667 | -34.5374 | -19.6369
11.8728 | -34.5378 | -19.6457
11.8789 | -34.5382 | -19.6546
11.885 | -34.5387 | -19.6635
11.8911 | -34.5391 | -19.6724
11.8972 | -34.5396 | -19.6812
11.9033 | -34.54 | -19.69
11.9094 | -34.5404 | -19.6988
11.9155 | -34.5408 | -19.7077
11.9216 | -34.5413 | -19.7163
11.9277 | -34.5417 | -19.7251
11.9338 | -34.5421 | -19.734
11.9399 | -34.5426 | -19.743
11.946 | -34.543 | -19.752
11.9521 | -34.5434 | -19.7608
11.9582 | -34.5439 | -19.7697
11.9644 | -34.5443 | -19.7784
11.9705 | -34.5448 | -19.7872
11.9766 | -34.5453 | -19.796
11.9827 | -34.5457 | -19.8049
11.9888 | -34.5462 | -19.8136
11.9949 | -34.5466 | -19.8222
12.001 | -34.5471 | -19.831
12.0071 | -34.5475 | -19.8399
12.0132 | -34.548 | -19.8488
12.0193 | -34.5484 | -19.8577
12.0254 | -34.5488 | -19.8667
12.0315 | -34.5492 | -19.8754
12.0376 | -34.5497 | -19.8842
12.0437 | -34.5501 | -19.8929
12.0498 | -34.5505 | -19.9019
12.0559 | -34.5509 | -19.9106
12.062 | -34.5514 | -19.9191
12.0681 | -34.5519 | -19.928
12.0742 | -34.5524 | -19.9367
12.0803 | -34.5528 | -19.9457
12.0864 | -34.5532 | -19.9545
12.0925 | -34.5537 | -19.9633
12.0986 | -34.5542 | -19.9719
12.1047 | -34.5546 | -19.9808
12.1109 | -34.5551 | -19.9895
12.117 | -34.5555 | -19.9982
12.1231 | -34.556 | -20.007
12.1292 | -34.5564 | -20.0156
12.1353 | -34.5568 | -20.0242
12.1414 | -34.5573 | -20.0331
12.1475 | -34.5577 | -20.0419
12.1536 | -34.5581 | -20.0508
12.1597 | -34.5586 | -20.0596
12.1658 | -34.5591 | -20.0684
12.1719 | -34.5595 | -20.0769
12.178 | -34.5599 | -20.0858
12.1841 | -34.5604 | -20.0946
12.1902 | -34.5608 | -20.1034
12.1963 | -34.5613 | -20.1122
12.2024 | -34.5618 | -20.1207
12.2085 | -34.5622 | -20.1294
12.2146 | -34.5627 | -20.1383
12.2207 | -34.5631 | -20.147
12.2268 | -34.5636 | -20.1557
12.2329 | -34.564 | -20.1645
12.239 | -34.5645 | -20.173
12.2451 | -34.565 | -20.1818
12.2513 | -34.5654 | -20.1908
12.2574 | -34.5658 | -20.1995
12.2635 | -34.5663 | -20.2083
12.2696 | -34.5667 | -20.217
12.2757 | -34.5672 | -20.2254
12.2818 | -34.5676 | -20.2342
12.2879 | -34.5681 | -20.243
12.294 | -34.5685 | -20.2517
12.3001 | -34.569 | -20.2604
12.3062 | -34.5694 | -20.2692
12.3123 | -34.5699 | -20.2776
12.3184 | -34.5703 | -20.2863
12.3245 | -34.5708 | -20.2948
12.3306 | -34.5712 | -20.3035
12.3367 | -34.5717 | -20.3123
12.3428 | -34.5722 | -20.321
12.3489 | -34.5727 | -20.3294
12.355 | -34.5731 | -20.3383
12.3611 | -34.5735 | -20.3472
12.3672 | -34.574 | -20.3558
12.3733 | -34.5744 | -20.3646
12.3794 | -34.5749 | -20.373
12.3855 | -34.5753 | -20.3816
12.3916 | -34.5758 | -20.3903
12.3978 | -34.5763 | -20.3991
12.4039 | -34.5767 | -20.4079
12.41 | -34.5771 | -20.4165
12.4161 | -34.5776 | -20.4248
12.4222 | -34.5781 | -20.4334
12.4283 | -34.5785 | -20.4421
12.4344 | -34.579 | -20.4508
12.4405 | -34.5794 | -20.4595
12.4466 | -34.5799 | -20.468
12.4527 | -34.5803 | -20.4763
12.4588 | -34.5808 | -20.485
12.4649 | -34.5812 | -20.4936
12.471 | -34.5817 | -20.5025
12.4771 | -34.5821 | -20.511
12.4832 | -34.5825 | -20.5195
12.4893 | -34.583 | -20.5279
12.4954 | -34.5835 | -20.5363
12.5015 | -34.5839 | -20.5449
12.5076 | -34.5844 | -20.5537
12.5137 | -34.5848 | -20.5623
12.5198 | -34.5853 | -20.5708
12.5259 | -34.5857 | -20.5794
12.532 | -34.5862 | -20.5879
12.5382 | -34.5866 | -20.5965
12.5443 | -34.5871 | -20.6051
12.5504 | -34.5875 | -20.6136
12.5565 | -34.588 | -20.6221
12.5626 | -34.5885 | -20.6309
12.5687 | -34.5889 | -20.6395
12.5748 | -34.5894 | -20.6483
12.5809 | -34.5898 | -20.6568
12.587 | -34.5903 | -20.6653
12.5931 | -34.5908 | -20.6736
12.5992 | -34.5913 | -20.6821
12.6053 | -34.5917 | -20.6906
12.6114 | -34.5922 | -20.6991
12.6175 | -34.5926 | -20.7078
12.6236 | -34.593 | -20.7164
12.6297 | -34.5935 | -20.7247
12.6358 | -34.594 | -20.7331
12.6419 | -34.5944 | -20.7416
12.648 | -34.5949 | -20.7501
12.6541 | -34.5953 | -20.7588
12.6602 | -34.5958 | -20.767
12.6663 | -34.5963 | -20.7756
12.6724 | -34.5967 | -20.784
12.6785 | -34.5972 | -20.7926
12.6847 | -34.5977 | -20.801
12.6908 | -34.5982 | -20.8097
12.6969 | -34.5986 | -20.8179
12.703 | -34.5991 | -20.8264
12.7091 | -34.5995 | -20.835
12.7152 | -34.6 | -20.8434
12.7213 | -34.6004 | -20.8521
12.7274 | -34.6009 | -20.8607
12.7335 | -34.6014 | -20.8688
12.7396 | -34.6018 | -20.877
12.7457 | -34.6023 | -20.8857
12.7518 | -34.6027 | -20.894
12.7579 | -34.6031 | -20.9027
12.764 | -34.6036 | -20.9113
12.7701 | -34.604 | -20.9198
12.7762 | -34.6045 | -20.9282
12.7823 | -34.605 | -20.9367
12.7884 | -34.6055 | -20.9452
12.7945 | -34.6059 | -20.9538
12.8006 | -34.6064 | -20.9619
12.8067 | -34.6069 | -20.9703
12.8128 | -34.6073 | -20.9786
12.8189 | -34.6078 | -20.9872
12.8251 | -34.6082 | -20.9957
12.8312 | -34.6087 | -21.0042
12.8373 | -34.6091 | -21.0125
12.8434 | -34.6096 | -21.0211
12.8495 | -34.61 | -21.0293
12.8556 | -34.6104 | -21.0376
12.8617 | -34.6109 | -21.046
12.8678 | -34.6114 | -21.0546
12.8739 | -34.6118 | -21.0629
12.88 | -34.6123 | -21.0712
12.8861 | -34.6127 | -21.0796
12.8922 | -34.6132 | -21.0883
12.8983 | -34.6137 | -21.0966
12.9044 | -34.6141 | -21.1048
12.9105 | -34.6146 | -21.1127
12.9166 | -34.615 | -21.1212
12.9227 | -34.6155 | -21.1295
12.9288 | -34.6159 | -21.1379
12.9349 | -34.6163 | -21.1464
12.941 | -34.6168 | -21.1547
12.9471 | -34.6173 | -21.1631
12.9532 | -34.6178 | -21.1715
12.9593 | -34.6183 | -21.18
12.9654 | -34.6187 | -21.1885
12.9716 | -34.6192 | -21.1968
12.9777 | -34.6197 | -21.2051
12.9838 | -34.6201 | -21.2133
12.9899 | -34.6206 | -21.2218
12.996 | -34.621 | -21.2301
13.0021 | -34.6215 | -21.2385
13.0082 | -34.622 | -21.2468
13.0143 | -34.6225 | -21.2546
13.0204 | -34.6229 | -21.2628
13.0265 | -34.6234 | -21.2712
13.0326 | -34.6238 | -21.2798
13.0387 | -34.6243 | -21.2884
13.0448 | -34.6247 | -21.2968
13.0509 | -34.6252 | -21.3051
13.057 | -34.6257 | -21.3135
13.0631 | -34.6261 | -21.3217
13.0692 | -34.6266 | -21.3299
13.0753 | -34.6271 | -21.3385
13.0814 | -34.6275 | -21.3472
13.0875 | -34.6279 | -21.3554
13.0936 | -34.6284 | -21.364
13.0997 | -34.6288 | -21.3724
13.1058 | -34.6293 | -21.3804
13.112 | -34.6298 | -21.3888
13.1181 | -34.6302 | -21.3965
13.1242 | -34.6307 | -21.4048
13.1303 | -34.6311 | -21.4131
13.1364 | -34.6316 | -21.4216
13.1425 | -34.632 | -21.4298
13.1486 | -34.6324 | -21.4382
13.1547 | -34.6329 | -21.4465
13.1608 | -34.6334 | -21.4547
13.1669 | -34.6338 | -21.4629
13.173 | -34.6343 | -21.4714
13.1791 | -34.6347 | -21.4799
13.1852 | -34.6351 | -21.4881
13.1913 | -34.6356 | -21.4965
13.1974 | -34.636 | -21.5049
13.2035 | -34.6364 | -21.5133
13.2096 | -34.6369 | -21.5216
13.2157 | -34.6374 | -21.5299
13.2218 | -34.6379 | -21.5382
13.2279 | -34.6383 | -21.5461
13.234 | -34.6388 | -21.5543
13.2401 | -34.6392 | -21.5626
13.2462 | -34.6397 | -21.5709
13.2524 | -34.6401 | -21.5793
13.2585 | -34.6406 | -21.5876
13.2646 | -34.641 | -21.5959
13.2707 | -34.6415 | -21.6042
13.2768 | -34.642 | -21.6128
13.2829 | -34.6424 | -21.621
13.289 | -34.6429 | -21.6291
13.2951 | -34.6433 | -21.6371
13.3012 | -34.6438 | -21.6455
13.3073 | -34.6443 | -21.6538
13.3134 | -34.6447 | -21.6623
13.3195 | -34.6452 | -21.6703
13.3256 | -34.6457 | -21.6786
13.3317 | -34.6461 | -21.6871
13.3378 | -34.6466 | -21.6952
13.3439 | -34.647 | -21.7036
13.35 | -34.6475 | -21.7121
13.3561 | -34.6479 | -21.7202
13.3622 | -34.6484 | -21.7283
13.3683 | -34.6489 | -21.736
13.3744 | -34.6493 | -21.7441
13.3805 | -34.6498 | -21.7524
13.3866 | -34.6502 | -21.7607
13.3927 | -34.6506 | -21.7689
13.3989 | -34.651 | -21.7771
13.405 | -34.6515 | -21.7855
13.4111 | -34.6519 | -21.7936
13.4172 | -34.6524 | -21.8019
13.4233 | -34.6529 | -21.8104
13.4294 | -34.6533 | -21.8189
13.4355 | -34.6537 | -21.8267
13.4416 | -34.6542 | -21.8349
13.4477 | -34.6546 | -21.8429
13.4538 | -34.6551 | -21.8514
13.4599 | -34.6555 | -21.8597
13.466 | -34.656 | -21.8679
13.4721 | -34.6564 | -21.8759
13.4782 | -34.6568 | -21.8842
13.4843 | -34.6573 | -21.8925
13.4904 | -34.6578 | -21.9005
13.4965 | -34.6583 | -21.9087
13.5026 | -34.6587 | -21.9169
13.5087 | -34.6592 | -21.9251
13.5148 | -34.6596 | -21.9336
13.5209 | -34.6601 | -21.9419
13.527 | -34.6605 | -21.9502
13.5331 | -34.661 | -21.9582
13.5393 | -34.6614 | -21.9663
13.5454 | -34.6618 | -21.9746
13.5515 | -34.6623 | -21.983
13.5576 | -34.6627 | -21.9915
13.5637 | -34.6632 | -21.9996
13.5698 | -34.6636 | -22.0078
13.5759 | -34.6641 | -22.0157
13.582 | -34.6645 | -22.024
13.5881 | -34.665 | -22.0321
13.5942 | -34.6655 | -22.0405
13.6003 | -34.6659 | -22.0487
13.6064 | -34.6664 | -22.057
13.6125 | -34.6668 | -22.0652
13.6186 | -34.6673 | -22.0733
13.6247 | -34.6677 | -22.0816
13.6308 | -34.6682 | -22.09
13.6369 | -34.6686 | -22.0981
13.643 | -34.6691 | -22.1063
13.6491 | -34.6695 | -22.1146
13.6552 | -34.6699 | -22.1227
13.6613 | -34.6704 | -22.1308
13.6674 | -34.6708 | -22.1388
13.6735 | -34.6713 | -22.1471
13.6796 | -34.6717 | -22.1553
13.6858 | -34.6722 | -22.1634
13.6919 | -34.6727 | -22.1718
13.698 | -34.6731 | -22.18
13.7041 | -34.6735 | -22.1885
13.7102 | -34.674 | -22.1968
13.7163 | -34.6745 | -22.2048
13.7224 | -34.675 | -22.2129
13.7285 | -34.6754 | -22.2211
13.7346 | -34.6759 | -22.2293
13.7407 | -34.6763 | -22.2376
13.7468 | -34.6767 | -22.2456
13.7529 | -34.6772 | -22.2538
13.759 | -34.6776 | -22.2619
13.7651 | -34.6781 | -22.2701
13.7712 | -34.6785 | -22.2785
13.7773 | -34.679 | -22.2866
13.7834 | -34.6795 | -22.2946
13.7895 | -34.6799 | -22.3026
13.7956 | -34.6803 | -22.311
13.8017 | -34.6808 | -22.3191
13.8078 | -34.6812 | -22.3272
13.8139 | -34.6817 | -22.3352
13.82 | -34.6821 | -22.3435
13.8262 | -34.6826 | -22.3518
13.8323 | -34.683 | -22.3601
13.8384 | -34.6835 | -22.3682
13.8445 | -34.684 | -22.3764
13.8506 | -34.6844 | -22.3846
13.8567 | -34.6848 | -22.393
13.8628 | -34.6853 | -22.4011
13.8689 | -34.6857 | -22.4093
13.875 | -34.6862 | -22.4176
13.8811 | -34.6866 | -22.4259
13.8872 | -34.687 | -22.4341
13.8933 | -34.6875 | -22.4421
13.8994 | -34.6879 | -22.4503
13.9055 | -34.6884 | -22.4586
13.9116 | -34.6889 | -22.4668
13.9177 | -34.6893 | -22.475
13.9238 | -34.6897 | -22.4832
13.9299 | -34.6902 | -22.4912
13.936 | -34.6906 | -22.4995
13.9421 | -34.6911 | -22.5077
13.9482 | -34.6915 | -22.5159
13.9543 | -34.6919 | -22.5242
13.9604 | -34.6924 | -22.5325
13.9665 | -34.6929 | -22.5405
13.9727 | -34.6933 | -22.5489
13.9788 | -34.6938 | -22.5568
13.9849 | -34.6943 | -22.565
13.991 | -34.6947 | -22.5731
13.9971 | -34.6952 | -22.581
14.0032 | -34.6956 | -22.5893
14.0093 | -34.696 | -22.5975
14.0154 | -34.6965 | -22.6055
14.0215 | -34.6969 | -22.6136
14.0276 | -34.6974 | -22.6218
14.0337 | -34.6978 | -22.6299
14.0398 | -34.6983 | -22.638
14.0459 | -34.6988 | -22.6461
14.052 | -34.6992 | -22.6545
14.0581 | -34.6996 | -22.6627
14.0642 | -34.7001 | -22.6709
14.0703 | -34.7005 | -22.6791
14.0764 | -34.7009 | -22.6872
14.0825 | -34.7014 | -22.6955
14.0886 | -34.7018 | -22.7037
14.0947 | -34.7023 | -22.7118
14.1008 | -34.7027 | -22.72
14.1069 | -34.7032 | -22.7281
14.1131 | -34.7036 | -22.7365
14.1192 | -34.7041 | -22.7445
14.1253 | -34.7045 | -22.7525
14.1314 | -34.705 | -22.7606
14.1375 | -34.7055 | -22.7687
14.1436 | -34.7059 | -22.7768
14.1497 | -34.7064 | -22.7848
14.1558 | -34.7068 | -22.793
14.1619 | -34.7072 | -22.801
14.168 | -34.7077 | -22.8093
14.1741 | -34.7082 | -22.8175
14.1802 | -34.7086 | -22.8257
14.1863 | -34.709 | -22.834
14.1924 | -34.7095 | -22.8421
14.1985 | -34.71 | -22.8502
14.2046 | -34.7104 | -22.8585
14.2107 | -34.7109 | -22.8665
14.2168 | -34.7113 | -22.8748
14.2229 | -34.7118 | -22.8829
14.229 | -34.7122 | -22.8913
14.2351 | -34.7126 | -22.8995
14.2412 | -34.7131 | -22.9078
14.2473 | -34.7135 | -22.9158
14.2534 | -34.714 | -22.9238
14.2596 | -34.7144 | -22.9319
14.2657 | -34.7148 | -22.9401
14.2718 | -34.7153 | -22.9482
14.2779 | -34.7157 | -22.9564
14.284 | -34.7162 | -22.9647
14.2901 | -34.7166 | -22.9728
14.2962 | -34.7171 | -22.9809
14.3023 | -34.7175 | -22.9891
14.3084 | -34.7179 | -22.9972
14.3145 | -34.7184 | -23.0055
14.3206 | -34.7188 | -23.0139
14.3267 | -34.7193 | -23.022
14.3328 | -34.7197 | -23.0301
14.3389 | -34.7202 | -23.0381
14.345 | -34.7206 | -23.0462
14.3511 | -34.721 | -23.0545
14.3572 | -34.7215 | -23.0628
14.3633 | -34.7219 | -23.071
14.3694 | -34.7224 | -23.0792
14.3755 | -34.7229 | -23.0873
14.3816 | -34.7233 | -23.0957
14.3877 | -34.7238 | -23.1036
14.3938 | -34.7242 | -23.112
14.4 | -34.7247 | -23.1201
14.4061 | -34.7251 | -23.1281
14.4122 | -34.7256 | -23.1363
14.4183 | -34.726 | -23.1443
14.4244 | -34.7265 | -23.1525
14.4305 | -34.7269 | -23.1607
14.4366 | -34.7274 | -23.1688
14.4427 | -34.7279 | -23.1771
14.4488 | -34.7283 | -23.1852
14.4549 | -34.7288 | -23.1933
14.461 | -34.7293 | -23.2013
14.4671 | -34.7297 | -23.2095
14.4732 | -34.7302 | -23.2178
14.4793 | -34.7306 | -23.2259
14.4854 | -34.7311 | -23.2338
14.4915 | -34.7316 | -23.2417
14.4976 | -34.732 | -23.2497
14.5037 | -34.7325 | -23.2577
14.5098 | -34.7329 | -23.2659
14.5159 | -34.7334 | -23.274
14.522 | -34.7338 | -23.2821
14.5281 | -34.7343 | -23.2902
14.5342 | -34.7347 | -23.2985
14.5403 | -34.7352 | -23.3066
14.5465 | -34.7356 | -23.3147
14.5526 | -34.736 | -23.323
14.5587 | -34.7365 | -23.3312
14.5648 | -34.737 | -23.3393
14.5709 | -34.7374 | -23.3474
14.577 | -34.7379 | -23.3556
14.5831 | -34.7383 | -23.3638
14.5892 | -34.7388 | -23.3718
14.5953 | -34.7392 | -23.3797
14.6014 | -34.7397 | -23.3878
14.6075 | -34.7401 | -23.3958
14.6136 | -34.7406 | -23.404
14.6197 | -34.741 | -23.4123
14.6258 | -34.7415 | -23.4204
14.6319 | -34.742 | -23.4286
14.638 | -34.7424 | -23.4368
14.6441 | -34.7428 | -23.445
14.6502 | -34.7433 | -23.4532
14.6563 | -34.7438 | -23.4614
14.6624 | -34.7442 | -23.4696
14.6685 | -34.7447 | -23.4775
14.6746 | -34.7452 | -23.4857
14.6807 | -34.7456 | -23.4938
14.6869 | -34.7461 | -23.5018
14.693 | -34.7465 | -23.5099
14.6991 | -34.747 | -23.518
14.7052 | -34.7475 | -23.5259
14.7113 | -34.748 | -23.5339
14.7174 | -34.7484 | -23.5421
14.7235 | -34.7489 | -23.5501
14.7296 | -34.7493 | -23.5581
14.7357 | -34.7498 | -23.5661
14.7418 | -34.7502 | -23.574
14.7479 | -34.7507 | -23.582
14.754 | -34.7511 | -23.5902
14.7601 | -34.7516 | -23.5982
14.7662 | -34.752 | -23.6062
14.7723 | -34.7525 | -23.614
14.7784 | -34.753 | -23.6221
14.7845 | -34.7534 | -23.6303
14.7906 | -34.7539 | -23.6383
14.7967 | -34.7543 | -23.6462
14.8028 | -34.7548 | -23.6542
14.8089 | -34.7553 | -23.6622
14.815 | -34.7557 | -23.6702
14.8211 | -34.7562 | -23.6783
14.8272 | -34.7566 | -23.6862
14.8334 | -34.7571 | -23.6945
14.8395 | -34.7576 | -23.7026
14.8456 | -34.758 | -23.7106
14.8517 | -34.7585 | -23.7187
14.8578 | -34.7589 | -23.7267
14.8639 | -34.7594 | -23.7348
14.87 | -34.7599 | -23.7428
14.8761 | -34.7603 | -23.7507
14.8822 | -34.7608 | -23.759
14.8883 | -34.7612 | -23.7669
14.8944 | -34.7617 | -23.7749
14.9005 | -34.7622 | -23.7829
14.9066 | -34.7626 | -23.791
14.9127 | -34.7631 | -23.7992
14.9188 | -34.7635 | -23.8073
14.9249 | -34.764 | -23.8155
14.931 | -34.7645 | -23.8236
14.9371 | -34.765 | -23.8315
14.9432 | -34.7654 | -23.8395
14.9493 | -34.7659 | -23.8475
14.9554 | -34.7664 | -23.8556
14.9615 | -34.7668 | -23.8637
14.9676 | -34.7673 | -23.8718
14.9738 | -34.7677 | -23.8798
14.9799 | -34.7682 | -23.8878
14.986 | -34.7686 | -23.8959
14.9921 | -34.7691 | -23.904
14.9982 | -34.7695 | -23.9119
15.0043 | -34.77 | -23.9199
15.0104 | -34.7705 | -23.9281
15.0165 | -34.7709 | -23.9364
15.0226 | -34.7714 | -23.9443
15.0287 | -34.7719 | -23.9523
15.0348 | -34.7724 | -23.9603
15.0409 | -34.7728 | -23.9685
15.047 | -34.7733 | -23.9768
15.0531 | -34.7738 | -23.9847
15.0592 | -34.7742 | -23.9927
15.0653 | -34.7747 | -24.0009
15.0714 | -34.7751 | -24.009
15.0775 | -34.7756 | -24.0169
15.0836 | -34.7761 | -24.0248
15.0897 | -34.7766 | -24.0326
15.0958 | -34.777 | -24.0406
15.1019 | -34.7775 | -24.0487
15.108 | -34.778 | -24.0566
15.1141 | -34.7785 | -24.0647
15.1203 | -34.7789 | -24.0728
15.1264 | -34.7794 | -24.0809
15.1325 | -34.7799 | -24.0891
15.1386 | -34.7803 | -24.0972
15.1447 | -34.7808 | -24.1052
15.1508 | -34.7813 | -24.1131
15.1569 | -34.7818 | -24.121
15.163 | -34.7823 | -24.1288
15.1691 | -34.7827 | -24.1368
15.1752 | -34.7832 | -24.1448
15.1813 | -34.7836 | -24.1528
15.1874 | -34.7841 | -24.1607
15.1935 | -34.7846 | -24.1685
15.1996 | -34.7851 | -24.1764
15.2057 | -34.7855 | -24.1843
15.2118 | -34.786 | -24.1923
15.2179 | -34.7865 | -24.2003
15.224 | -34.7869 | -24.2082
15.2301 | -34.7874 | -24.2158
15.2362 | -34.7878 | -24.224
15.2423 | -34.7883 | -24.2321
15.2484 | -34.7888 | -24.24
15.2545 | -34.7893 | -24.2481
15.2607 | -34.7898 | -24.2561
15.2668 | -34.7902 | -24.2636
15.2729 | -34.7907 | -24.2715
15.279 | -34.7912 | -24.2797
15.2851 | -34.7917 | -24.2876
15.2912 | -34.7922 | -24.2956
15.2973 | -34.7927 | -24.3037
15.3034 | -34.7931 | -24.3114
15.3095 | -34.7936 | -24.3194
15.3156 | -34.7941 | -24.3273
15.3217 | -34.7946 | -24.3353
15.3278 | -34.795 | -24.3435
15.3339 | -34.7955 | -24.3513
15.34 | -34.796 | -24.3591
15.3461 | -34.7965 | -24.3669
15.3522 | -34.7969 | -24.3749
15.3583 | -34.7974 | -24.3831
15.3644 | -34.7979 | -24.3908
15.3705 | -34.7984 | -24.3984
15.3766 | -34.7988 | -24.4063
15.3827 | -34.7993 | -24.4141
15.3888 | -34.7998 | -24.4221
15.3949 | -34.8003 | -24.4302
15.401 | -34.8007 | -24.4381
15.4072 | -34.8012 | -24.4458
15.4133 | -34.8017 | -24.4538
15.4194 | -34.8021 | -24.4616
15.4255 | -34.8026 | -24.4695
15.4316 | -34.8031 | -24.4775
15.4377 | -34.8035 | -24.4855
15.4438 | -34.804 | -24.4933
15.4499 | -34.8045 | -24.5011
15.456 | -34.805 | -24.5089
15.4621 | -34.8055 | -24.5166
15.4682 | -34.8059 | -24.5246
15.4743 | -34.8065 | -24.5323
15.4804 | -34.8069 | -24.54
15.4865 | -34.8074 | -24.548
15.4926 | -34.8079 | -24.556
15.4987 | -34.8084 | -24.5639
15.5048 | -34.8089 | -24.5717
15.5109 | -34.8093 | -24.5796
15.517 | -34.8098 | -24.5875
15.5231 | -34.8103 | -24.5953
15.5292 | -34.8107 | -24.6031
15.5353 | -34.8112 | -24.6112
15.5414 | -34.8116 | -24.6189
15.5476 | -34.8121 | -24.6268
15.5537 | -34.8126 | -24.6345
15.5598 | -34.8131 | -24.6427
15.5659 | -34.8135 | -24.6506
15.572 | -34.814 | -24.6585
15.5781 | -34.8145 | -24.6664
15.5842 | -34.815 | -24.6743
15.5903 | -34.8155 | -24.682
15.5964 | -34.816 | -24.6897
15.6025 | -34.8164 | -24.6973
15.6086 | -34.817 | -24.7051
15.6147 | -34.8174 | -24.713
15.6208 | -34.818 | -24.7205
15.6269 | -34.8185 | -24.7283
15.633 | -34.8189 | -24.7363
15.6391 | -34.8194 | -24.7441
15.6452 | -34.8199 | -24.752
15.6513 | -34.8204 | -24.7596
15.6574 | -34.8209 | -24.7676
15.6635 | -34.8214 | -24.7754
15.6696 | -34.8218 | -24.7832
15.6757 | -34.8223 | -24.791
15.6818 | -34.8228 | -24.7988
15.688 | -34.8233 | -24.8065
15.6941 | -34.8237 | -24.8144
15.7002 | -34.8242 | -24.8221
15.7063 | -34.8247 | -24.8299
15.7124 | -34.8252 | -24.8377
15.7185 | -34.8256 | -24.8454
15.7246 | -34.8262 | -24.8532
15.7307 | -34.8267 | -24.8608
15.7368 | -34.8271 | -24.8687
15.7429 | -34.8276 | -24.8764
15.749 | -34.8281 | -24.8843
15.7551 | -34.8286 | -24.8922
15.7612 | -34.8291 | -24.8999
15.7673 | -34.8296 | -24.9076
15.7734 | -34.8301 | -24.9154
15.7795 | -34.8306 | -24.9231
15.7856 | -34.8311 | -24.9307
15.7917 | -34.8316 | -24.9383
15.7978 | -34.832 | -24.9463
15.8039 | -34.8325 | -24.954
15.81 | -34.833 | -24.9619
15.8161 | -34.8335 | -24.9699
15.8222 | -34.834 | -24.9775
15.8283 | -34.8345 | -24.985
15.8345 | -34.835 | -24.9928
15.8406 | -34.8355 | -25.0005
15.8467 | -34.836 | -25.0084
15.8528 | -34.8365 | -25.0164
15.8589 | -34.8369 | -25.0243
15.865 | -34.8375 | -25.0317
15.8711 | -34.8379 | -25.0396
15.8772 | -34.8384 | -25.0471
15.8833 | -34.8389 | -25.0549
15.8894 | -34.8393 | -25.0627
15.8955 | -34.8398 | -25.0704
15.9016 | -34.8402 | -25.0777
15.9077 | -34.8407 | -25.0855
15.9138 | -34.8412 | -25.0934
15.9199 | -34.8417 | -25.1008
15.926 | -34.8421 | -25.1084
15.9321 | -34.8426 | -25.116
15.9382 | -34.8431 | -25.1235
15.9443 | -34.8436 | -25.1311
15.9504 | -34.8441 | -25.1388
15.9565 | -34.8446 | -25.1463
15.9626 | -34.845 | -25.1541
15.9687 | -34.8455 | -25.1616
15.9749 | -34.846 | -25.1692
15.981 | -34.8465 | -25.1769
15.9871 | -34.8469 | -25.1847
15.9932 | -34.8474 | -25.1922
15.9993 | -34.8479 | -25.2
16.0054 | -34.8484 | -25.2073
16.0115 | -34.8489 | -25.215
16.0176 | -34.8494 | -25.2228
16.0237 | -34.8499 | -25.2307
16.0298 | -34.8503 | -25.2382
16.0359 | -34.8508 | -25.2458
16.042 | -34.8513 | -25.2534
16.0481 | -34.8518 | -25.261
16.0542 | -34.8523 | -25.269
16.0603 | -34.8528 | -25.2765
16.0664 | -34.8532 | -25.2842
16.0725 | -34.8537 | -25.2916
16.0786 | -34.8542 | -25.2992
16.0847 | -34.8547 | -25.307
16.0908 | -34.8552 | -25.3146
16.0969 | -34.8557 | -25.3224
16.103 | -34.8562 | -25.3302
16.1091 | -34.8567 | -25.3376
16.1152 | -34.8572 | -25.3454
16.1214 | -34.8577 | -25.3531
16.1275 | -34.8582 | -25.3607
16.1336 | -34.8586 | -25.3683
16.1397 | -34.8591 | -25.3761
16.1458 | -34.8596 | -25.3836
16.1519 | -34.8601 | -25.3916
16.158 | -34.8606 | -25.3992
16.1641 | -34.8611 | -25.4068
16.1702 | -34.8615 | -25.4145
16.1763 | -34.862 | -25.4221
16.1824 | -34.8625 | -25.4296
16.1885 | -34.863 | -25.4372
16.1946 | -34.8635 | -25.4448
16.2007 | -34.8639 | -25.4526
16.2068 | -34.8644 | -25.4601
16.2129 | -34.8649 | -25.4678
16.219 | -34.8655 | -25.4752
16.2251 | -34.8659 | -25.4828
16.2312 | -34.8664 | -25.4906
16.2373 | -34.8669 | -25.4981
16.2434 | -34.8674 | -25.5058
16.2495 | -34.8679 | -25.5133
16.2556 | -34.8684 | -25.521
16.2618 | -34.8689 | -25.5288
16.2679 | -34.8694 | -25.5365
16.274 | -34.8698 | -25.5441
16.2801 | -34.8703 | -25.5519
16.2862 | -34.8708 | -25.5597
16.2923 | -34.8713 | -25.5672
16.2984 | -34.8717 | -25.5749
16.3045 | -34.8722 | -25.5824
16.3106 | -34.8727 | -25.5901
16.3167 | -34.8732 | -25.5977
16.3228 | -34.8737 | -25.6053
16.3289 | -34.8741 | -25.6128
16.335 | -34.8747 | -25.6203
16.3411 | -34.8752 | -25.6279
16.3472 | -34.8756 | -25.6357
16.3533 | -34.8761 | -25.6434
16.3594 | -34.8766 | -25.6505
16.3655 | -34.8772 | -25.6582
16.3716 | -34.8776 | -25.6657
16.3777 | -34.8781 | -25.6734
16.3838 | -34.8786 | -25.6811
16.3899 | -34.8791 | -25.6886
16.396 | -34.8796 | -25.6962
16.4021 | -34.88 | -25.7038
16.4083 | -34.8805 | -25.7113
16.4144 | -34.881 | -25.7189
16.4205 | -34.8815 | -25.7265
16.4266 | -34.882 | -25.7336
16.4327 | -34.8825 | -25.7411
16.4388 | -34.883 | -25.7485
16.4449 | -34.8835 | -25.7561
16.451 | -34.8839 | -25.7637
16.4571 | -34.8844 | -25.7712
16.4632 | -34.8848 | -25.7785
16.4693 | -34.8852 | -25.7862
16.4754 | -34.8857 | -25.7938
16.4815 | -34.8862 | -25.8011
16.4876 | -34.8867 | -25.8085
16.4937 | -34.8871 | -25.816
16.4998 | -34.8877 | -25.8232
16.5059 | -34.8882 | -25.8307
16.512 | -34.8887 | -25.8385
16.5181 | -34.8892 | -25.8463
16.5242 | -34.8897 | -25.854
16.5303 | -34.8901 | -25.8615
16.5364 | -34.8906 | -25.8691
16.5425 | -34.8911 | -25.8766
16.5487 | -34.8915 | -25.8842
16.5548 | -34.892 | -25.8918
16.5609 | -34.8926 | -25.8995
16.567 | -34.8931 | -25.9072
16.5731 | -34.8936 | -25.9149
16.5792 | -34.8941 | -25.9225
16.5853 | -34.8946 | -25.9302
16.5914 | -34.895 | -25.9377
16.5975 | -34.8955 | -25.9455
16.6036 | -34.896 | -25.9529
16.6097 | -34.8964 | -25.9607
16.6158 | -34.8969 | -25.9681
16.6219 | -34.8974 | -25.9758
16.628 | -34.8979 | -25.9835
16.6341 | -34.8983 | -25.9909
16.6402 | -34.8988 | -25.9986
16.6463 | -34.8993 | -26.0062
16.6524 | -34.8998 | -26.0137
16.6585 | -34.9003 | -26.0214
16.6646 | -34.9008 | -26.0288
16.6707 | -34.9013 | -26.0365
16.6768 | -34.9018 | -26.0442
16.6829 | -34.9023 | -26.0516
16.689 | -34.9027 | -26.0592
16.6952 | -34.9032 | -26.0669
16.7013 | -34.9037 | -26.0743
16.7074 | -34.9042 | -26.0816
16.7135 | -34.9047 | -26.0893
16.7196 | -34.9051 | -26.0969
16.7257 | -34.9056 | -26.1043
16.7318 | -34.9061 | -26.1115
16.7379 | -34.9066 | -26.1188
16.744 | -34.9071 | -26.1261
16.7501 | -34.9075 | -26.1335
16.7562 | -34.908 | -26.1411
16.7623 | -34.9084 | -26.1488
16.7684 | -34.9089 | -26.1566
16.7745 | -34.9094 | -26.1643
16.7806 | -34.9099 | -26.1715
16.7867 | -34.9104 | -26.1788
16.7928 | -34.9109 | -26.1862
16.7989 | -34.9113 | -26.1936
16.805 | -34.9118 | -26.2014
16.8111 | -34.9123 | -26.2087
16.8172 | -34.9128 | -26.2163
16.8233 | -34.9133 | -26.2238
16.8294 | -34.9138 | -26.2313
16.8356 | -34.9143 | -26.2386
16.8417 | -34.9147 | -26.2461
16.8478 | -34.9152 | -26.2533
16.8539 | -34.9157 | -26.2605
16.86 | -34.9161 | -26.268
16.8661 | -34.9166 | -26.2752
16.8722 | -34.917 | -26.2828
16.8783 | -34.9176 | -26.2903
16.8844 | -34.9181 | -26.2976
16.8905 | -34.9185 | -26.3053
16.8966 | -34.919 | -26.3126
16.9027 | -34.9195 | -26.3201
16.9088 | -34.92 | -26.3277
16.9149 | -34.9205 | -26.3354
16.921 | -34.921 | -26.3428
16.9271 | -34.9214 | -26.3503
16.9332 | -34.9219 | -26.3578
16.9393 | -34.9224 | -26.3654
16.9454 | -34.9228 | -26.3728
16.9515 | -34.9233 | -26.3805
16.9576 | -34.9238 | -26.3879
16.9637 | -34.9243 | -26.3954
16.9698 | -34.9248 | -26.4027
16.9759 | -34.9252 | -26.4101
16.9821 | -34.9257 | -26.4175
16.9882 | -34.9262 | -26.4251
16.9943 | -34.9267 | -26.4325
17.0004 | -34.9272 | -26.4397
17.0065 | -34.9277 | -26.4471
17.0126 | -34.9282 | -26.4545
17.0187 | -34.9286 | -26.4622
17.0248 | -34.9291 | -26.4697
17.0309 | -34.9296 | -26.4773
17.037 | -34.93 | -26.4848
17.0431 | -34.9305 | -26.4923
17.0492 | -34.931 | -26.4999
17.0553 | -34.9314 | -26.5075
17.0614 | -34.9319 | -26.515
17.0675 | -34.9324 | -26.5224
17.0736 | -34.9329 | -26.5298
17.0797 | -34.9334 | -26.5372
17.0858 | -34.9339 | -26.5447
17.0919 | -34.9344 | -26.5522
17.098 | -34.9349 | -26.5595
17.1041 | -34.9353 | -26.5669
17.1102 | -34.9358 | -26.5742
17.1163 | -34.9363 | -26.5816
17.1225 | -34.9368 | -26.5888
17.1286 | -34.9373 | -26.5962
17.1347 | -34.9378 | -26.6038
17.1408 | -34.9383 | -26.6114
17.1469 | -34.9388 | -26.6188
17.153 | -34.9393 | -26.6262
17.1591 | -34.9397 | -26.6336
17.1652 | -34.9402 | -26.6407
17.1713 | -34.9407 | -26.648
17.1774 | -34.9412 | -26.6556
17.1835 | -34.9416 | -26.6632
17.1896 | -34.9421 | -26.6709
17.1957 | -34.9426 | -26.6783
17.2018 | -34.943 | -26.6856
17.2079 | -34.9435 | -26.693
17.214 | -34.944 | -26.7005
17.2201 | -34.9445 | -26.708
17.2262 | -34.9449 | -26.7157
17.2323 | -34.9454 | -26.7232
17.2384 | -34.9459 | -26.7306
17.2445 | -34.9464 | -26.738
17.2506 | -34.9469 | -26.7455
17.2567 | -34.9474 | -26.753
17.2628 | -34.9478 | -26.7604
17.269 | -34.9483 | -26.7676
17.2751 | -34.9488 | -26.7751
17.2812 | -34.9492 | -26.7826
17.2873 | -34.9497 | -26.79
17.2934 | -34.9502 | -26.7976
17.2995 | -34.9507 | -26.8051
17.3056 | -34.9511 | -26.8122
17.3117 | -34.9516 | -26.8199
17.3178 | -34.9521 | -26.8271
17.3239 | -34.9525 | -26.8344
17.33 | -34.953 | -26.8418
17.3361 | -34.9535 | -26.8488
17.3422 | -34.9539 | -26.856
17.3483 | -34.9544 | -26.8633
17.3544 | -34.9549 | -26.8708
17.3605 | -34.9554 | -26.8781
17.3666 | -34.9559 | -26.8854
17.3727 | -34.9564 | -26.8927
17.3788 | -34.9569 | -26.9003
17.3849 | -34.9573 | -26.9076
17.391 | -34.9578 | -26.9151
17.3971 | -34.9583 | -26.9225
17.4032 | -34.9587 | -26.93
17.4094 | -34.9592 | -26.9374
17.4155 | -34.9597 | -26.945
17.4216 | -34.9602 | -26.9524
17.4277 | -34.9607 | -26.9602
17.4338 | -34.9612 | -26.9674
17.4399 | -34.9617 | -26.975
17.446 | -34.9622 | -26.9823
17.4521 | -34.9627 | -26.9898
17.4582 | -34.9632 | -26.9968
17.4643 | -34.9636 | -27.0039
17.4704 | -34.9641 | -27.0113
17.4765 | -34.9646 | -27.0185
17.4826 | -34.9651 | -27.026
17.4887 | -34.9656 | -27.0336
17.4948 | -34.9661 | -27.0407
17.5009 | -34.9666 | -27.0481
17.507 | -34.9671 | -27.0555
17.5131 | -34.9675 | -27.063
17.5192 | -34.968 | -27.0702
17.5253 | -34.9685 | -27.0774
17.5314 | -34.9689 | -27.0849
17.5375 | -34.9694 | -27.0923
17.5436 | -34.9699 | -27.0997
17.5497 | -34.9704 | -27.1071
17.5559 | -34.9709 | -27.1142
17.562 | -34.9714 | -27.1213
17.5681 | -34.9719 | -27.1289
17.5742 | -34.9724 | -27.1365
17.5803 | -34.9728 | -27.144
17.5864 | -34.9733 | -27.1517
17.5925 | -34.9738 | -27.1591
17.5986 | -34.9743 | -27.1663
17.6047 | -34.9748 | -27.1739
17.6108 | -34.9753 | -27.181
17.6169 | -34.9758 | -27.1884
17.623 | -34.9762 | -27.1957
17.6291 | -34.9767 | -27.2029
17.6352 | -34.9772 | -27.2103
17.6413 | -34.9777 | -27.2177
17.6474 | -34.9782 | -27.2251
17.6535 | -34.9786 | -27.2324
17.6596 | -34.9791 | -27.2399
17.6657 | -34.9796 | -27.2474
17.6718 | -34.9801 | -27.2548
17.6779 | -34.9806 | -27.2623
17.684 | -34.9811 | -27.2695
17.6901 | -34.9815 | -27.2769
17.6963 | -34.982 | -27.2844
17.7024 | -34.9825 | -27.2917
17.7085 | -34.983 | -27.2992
17.7146 | -34.9835 | -27.3062
17.7207 | -34.984 | -27.3134
17.7268 | -34.9845 | -27.3207
17.7329 | -34.985 | -27.3282
17.739 | -34.9854 | -27.3358
17.7451 | -34.9859 | -27.343
17.7512 | -34.9864 | -27.3504
17.7573 | -34.9869 | -27.3577
17.7634 | -34.9873 | -27.365
17.7695 | -34.9878 | -27.3722
17.7756 | -34.9883 | -27.3796
17.7817 | -34.9888 | -27.3871
17.7878 | -34.9893 | -27.3947
17.7939 | -34.9898 | -27.4018
17.8 | -34.9902 | -27.4094
17.8061 | -34.9907 | -27.4167
17.8122 | -34.9912 | -27.4241
17.8183 | -34.9917 | -27.4314
17.8244 | -34.9921 | -27.4389
17.8305 | -34.9926 | -27.4462
17.8366 | -34.9931 | -27.4534
17.8428 | -34.9935 | -27.4607
17.8489 | -34.994 | -27.4678
17.855 | -34.9945 | -27.4749
17.8611 | -34.9949 | -27.4821
17.8672 | -34.9954 | -27.4896
17.8733 | -34.9959 | -27.497
17.8794 | -34.9964 | -27.5043
17.8855 | -34.9968 | -27.5118
17.8916 | -34.9973 | -27.5195
17.8977 | -34.9977 | -27.527
17.9038 | -34.9982 | -27.5336
17.9099 | -34.9987 | -27.5409
17.916 | -34.9992 | -27.5483
17.9221 | -34.9997 | -27.5557
17.9282 | -35.0001 | -27.5629
17.9343 | -35.0006 | -27.5705
17.9404 | -35.0011 | -27.5776
17.9465 | -35.0016 | -27.5851
17.9526 | -35.002 | -27.5924
17.9587 | -35.0025 | -27.5999
17.9648 | -35.0029 | -27.607
17.9709 | -35.0034 | -27.6144
17.977 | -35.0039 | -27.6216
17.9832 | -35.0044 | -27.6289
17.9893 | -35.0049 | -27.6365
17.9954 | -35.0053 | -27.6439
18.0015 | -35.0058 | -27.6513
18.0076 | -35.0063 | -27.6589
18.0137 | -35.0068 | -27.6663
18.0198 | -35.0072 | -27.6735
18.0259 | -35.0077 | -27.6812
18.032 | -35.0082 | -27.6882
18.0381 | -35.0086 | -27.6955
18.0442 | -35.0092 | -27.7028
18.0503 | -35.0096 | -27.7103
18.0564 | -35.0101 | -27.7179
18.0625 | -35.0105 | -27.7254
18.0686 | -35.011 | -27.7328
18.0747 | -35.0114 | -27.7401
18.0808 | -35.0119 | -27.7473
18.0869 | -35.0123 | -27.7549
18.093 | -35.0128 | -27.7623
18.0991 | -35.0133 | -27.7698
18.1052 | -35.0137 | -27.7772
18.1113 | -35.0142 | -27.7845
18.1174 | -35.0146 | -27.7917
18.1236 | -35.0151 | -27.7991
18.1297 | -35.0156 | -27.8065
18.1358 | -35.0161 | -27.814
18.1419 | -35.0166 | -27.8213
18.148 | -35.0171 | -27.8284
18.1541 | -35.0175 | -27.8356
18.1602 | -35.018 | -27.843
18.1663 | -35.0184 | -27.8502
18.1724 | -35.0189 | -27.8576
18.1785 | -35.0194 | -27.8648
18.1846 | -35.0198 | -27.8719
18.1907 | -35.0203 | -27.8793
18.1968 | -35.0208 | -27.8869
18.2029 | -35.0212 | -27.8943
18.209 | -35.0217 | -27.9017
18.2151 | -35.0221 | -27.9091
18.2212 | -35.0226 | -27.9159
18.2273 | -35.0231 | -27.9233
18.2334 | -35.0236 | -27.9307
18.2395 | -35.0241 | -27.938
18.2456 | -35.0246 | -27.9453
18.2517 | -35.025 | -27.9522
18.2578 | -35.0255 | -27.9591
18.2639 | -35.026 | -27.9666
18.2701 | -35.0265 | -27.9739
18.2762 | -35.0269 | -27.9812
18.2823 | -35.0274 | -27.9885
18.2884 | -35.0279 | -27.9958
18.2945 | -35.0284 | -28.0029
18.3006 | -35.0288 | -28.0102
18.3067 | -35.0293 | -28.0174
18.3128 | -35.0298 | -28.0248
18.3189 | -35.0302 | -28.0318
18.325 | -35.0307 | -28.039
18.3311 | -35.0312 | -28.0463
18.3372 | -35.0316 | -28.0535
18.3433 | -35.0321 | -28.0609
18.3494 | -35.0326 | -28.0683
18.3555 | -35.033 | -28.0758
18.3616 | -35.0335 | -28.0831
18.3677 | -35.034 | -28.0902
18.3738 | -35.0345 | -28.0973
18.3799 | -35.035 | -28.1043
18.386 | -35.0354 | -28.1115
18.3921 | -35.0359 | -28.1187
18.3982 | -35.0364 | -28.1259
18.4043 | -35.0368 | -28.1331
18.4105 | -35.0372 | -28.1405
18.4166 | -35.0377 | -28.1479
18.4227 | -35.0382 | -28.1552
18.4288 | -35.0386 | -28.1626
18.4349 | -35.039 | -28.1698
18.441 | -35.0395 | -28.177
18.4471 | -35.04 | -28.1842
18.4532 | -35.0405 | -28.1914
18.4593 | -35.041 | -28.1985
18.4654 | -35.0414 | -28.2055
18.4715 | -35.0419 | -28.2126
18.4776 | -35.0424 | -28.2202
18.4837 | -35.0428 | -28.2275
18.4898 | -35.0433 | -28.2352
18.4959 | -35.0438 | -28.2426
18.502 | -35.0442 | -28.2501
18.5081 | -35.0447 | -28.2575
18.5142 | -35.0452 | -28.2646
18.5203 | -35.0456 | -28.2715
18.5264 | -35.0461 | -28.2786
18.5325 | -35.0466 | -28.2856
18.5386 | -35.047 | -28.293
18.5447 | -35.0475 | -28.3002
18.5508 | -35.048 | -28.3072
18.557 | -35.0484 | -28.3145
18.5631 | -35.0489 | -28.3218
18.5692 | -35.0493 | -28.3291
18.5753 | -35.0498 | -28.3361
18.5814 | -35.0503 | -28.3432
18.5875 | -35.0507 | -28.3503
18.5936 | -35.0512 | -28.3574
18.5997 | -35.0516 | -28.3644
18.6058 | -35.0521 | -28.3715
18.6119 | -35.0526 | -28.3785
18.618 | -35.0531 | -28.3855
18.6241 | -35.0535 | -28.393
18.6302 | -35.054 | -28.4002
18.6363 | -35.0544 | -28.4073
18.6424 | -35.0549 | -28.4147
18.6485 | -35.0553 | -28.4219
18.6546 | -35.0558 | -28.4291
18.6607 | -35.0563 | -28.4366
18.6668 | -35.0568 | -28.4439
18.6729 | -35.0573 | -28.451
18.679 | -35.0577 | -28.4582
18.6851 | -35.0582 | -28.4653
18.6912 | -35.0586 | -28.4721
18.6974 | -35.0591 | -28.4791
18.7035 | -35.0595 | -28.4867
18.7096 | -35.06 | -28.4941
18.7157 | -35.0605 | -28.5016
18.7218 | -35.061 | -28.5088
18.7279 | -35.0615 | -28.5156
18.734 | -35.0619 | -28.5229
18.7401 | -35.0624 | -28.5301
18.7462 | -35.0629 | -28.5367
18.7523 | -35.0634 | -28.544
18.7584 | -35.0639 | -28.5513
18.7645 | -35.0644 | -28.5584
18.7706 | -35.0648 | -28.5656
18.7767 | -35.0652 | -28.5727
18.7828 | -35.0657 | -28.5795
18.7889 | -35.0662 | -28.5868
18.795 | -35.0666 | -28.594
18.8011 | -35.0671 | -28.6013
18.8072 | -35.0676 | -28.6086
18.8133 | -35.0681 | -28.6157
18.8194 | -35.0685 | -28.6228
18.8255 | -35.069 | -28.63
18.8316 | -35.0695 | -28.6371
18.8377 | -35.0699 | -28.6443
18.8439 | -35.0704 | -28.6511
18.85 | -35.0709 | -28.6583
18.8561 | -35.0712 | -28.6651
18.8622 | -35.0718 | -28.6721
18.8683 | -35.0722 | -28.679
18.8744 | -35.0727 | -28.6862
18.8805 | -35.0732 | -28.6935
18.8866 | -35.0736 | -28.7008
18.8927 | -35.0741 | -28.7078
18.8988 | -35.0745 | -28.715
18.9049 | -35.075 | -28.7222
18.911 | -35.0754 | -28.7294
18.9171 | -35.0759 | -28.7366
18.9232 | -35.0763 | -28.7437
18.9293 | -35.0768 | -28.751
18.9354 | -35.0773 | -28.7582
18.9415 | -35.0778 | -28.7654
18.9476 | -35.0783 | -28.7723
18.9537 | -35.0787 | -28.7798
18.9598 | -35.0792 | -28.7867
18.9659 | -35.0797 | -28.7938
18.972 | -35.0802 | -28.801
18.9781 | -35.0806 | -28.8083
18.9843 | -35.081 | -28.8159
18.9904 | -35.0815 | -28.8233
18.9965 | -35.082 | -28.8302
19.0026 | -35.0825 | -28.8377
19.0087 | -35.0829 | -28.8449
19.0148 | -35.0833 | -28.8521
19.0209 | -35.0838 | -28.8593
19.027 | -35.0843 | -28.8667
19.0331 | -35.0847 | -28.8734
19.0392 | -35.0851 | -28.8806
19.0453 | -35.0856 | -28.8877
19.0514 | -35.0861 | -28.8947
19.0575 | -35.0865 | -28.9019
19.0636 | -35.087 | -28.9087
19.0697 | -35.0875 | -28.9154
19.0758 | -35.0879 | -28.9226
19.0819 | -35.0884 | -28.9297
19.088 | -35.0889 | -28.937
19.0941 | -35.0893 | -28.9444
19.1002 | -35.0898 | -28.9515
19.1063 | -35.0902 | -28.9587
19.1124 | -35.0907 | -28.9659
19.1185 | -35.0911 | -28.9727
19.1246 | -35.0916 | -28.9802
19.1308 | -35.092 | -28.9874
19.1369 | -35.0925 | -28.9946
19.143 | -35.093 | -29.002
19.1491 | -35.0934 | -29.0089
19.1552 | -35.0939 | -29.0159
19.1613 | -35.0943 | -29.0231
19.1674 | -35.0948 | -29.0303
19.1735 | -35.0953 | -29.0375
19.1796 | -35.0958 | -29.0444
19.1857 | -35.0963 | -29.0517
19.1918 | -35.0967 | -29.0588
19.1979 | -35.0971 | -29.0657
19.204 | -35.0975 | -29.0728
19.2101 | -35.098 | -29.0799
19.2162 | -35.0984 | -29.087
19.2223 | -35.0989 | -29.0941
19.2284 | -35.0994 | -29.1012
19.2345 | -35.0998 | -29.1086
19.2406 | -35.1002 | -29.1152
19.2467 | -35.1007 | -29.1222
19.2528 | -35.1012 | -29.1292
19.2589 | -35.1016 | -29.1365
19.265 | -35.1022 | -29.1436
19.2712 | -35.1026 | -29.1507
19.2773 | -35.1031 | -29.1576
19.2834 | -35.1036 | -29.1646
19.2895 | -35.104 | -29.1718
19.2956 | -35.1044 | -29.1789
19.3017 | -35.1048 | -29.1857
19.3078 | -35.1052 | -29.193
19.3139 | -35.1057 | -29.1999
19.32 | -35.1062 | -29.2066
19.3261 | -35.1065 | -29.2138
19.3322 | -35.107 | -29.2212
19.3383 | -35.1075 | -29.2283
19.3444 | -35.1079 | -29.2351
19.3505 | -35.1084 | -29.2424
19.3566 | -35.1088 | -29.2493
19.3627 | -35.1093 | -29.2562
19.3688 | -35.1097 | -29.2633
19.3749 | -35.1102 | -29.2707
19.381 | -35.1107 | -29.2775
19.3871 | -35.1111 | -29.2844
19.3932 | -35.1116 | -29.2912
19.3993 | -35.1121 | -29.2981
19.4054 | -35.1125 | -29.305
19.4115 | -35.113 | -29.3122
19.4177 | -35.1135 | -29.3187
19.4238 | -35.1139 | -29.3258
19.4299 | -35.1144 | -29.3328
19.436 | -35.1148 | -29.3399
19.4421 | -35.1153 | -29.347
19.4482 | -35.1157 | -29.3541
19.4543 | -35.1162 | -29.3606
19.4604 | -35.1166 | -29.3673
19.4665 | -35.1171 | -29.3744
19.4726 | -35.1175 | -29.3811
19.4787 | -35.1179 | -29.3881
19.4848 | -35.1183 | -29.3953
19.4909 | -35.1187 | -29.4019
19.497 | -35.1191 | -29.4091
19.5031 | -35.1196 | -29.416
19.5092 | -35.12 | -29.4229
19.5153 | -35.1205 | -29.4302
19.5214 | -35.121 | -29.4368
19.5275 | -35.1214 | -29.4435
19.5336 | -35.1218 | -29.4503
19.5397 | -35.1223 | -29.4572
19.5458 | -35.1227 | -29.4638
19.5519 | -35.1232 | -29.4704
19.5581 | -35.1236 | -29.4768
19.5642 | -35.1241 | -29.484
19.5703 | -35.1245 | -29.4909
19.5764 | -35.125 | -29.4981
19.5825 | -35.1254 | -29.5053
19.5886 | -35.1259 | -29.5122
19.5947 | -35.1265 | -29.5188
19.6008 | -35.1269 | -29.5255
19.6069 | -35.1273 | -29.5324
19.613 | -35.1278 | -29.5392
19.6191 | -35.1282 | -29.5461
19.6252 | -35.1287 | -29.5534
19.6313 | -35.1291 | -29.5601
19.6374 | -35.1295 | -29.5668
19.6435 | -35.1299 | -29.5736
19.6496 | -35.1304 | -29.5806
19.6557 | -35.1308 | -29.5877
19.6618 | -35.1313 | -29.5946
19.6679 | -35.1317 | -29.6015
19.674 | -35.1322 | -29.6084
19.6801 | -35.1326 | -29.6152
19.6862 | -35.1331 | -29.6218
19.6923 | -35.1335 | -29.6284
19.6984 | -35.1339 | -29.6348
19.7046 | -35.1344 | -29.6415
19.7107 | -35.1349 | -29.6488
19.7168 | -35.1354 | -29.6557
19.7229 | -35.1358 | -29.6632
19.729 | -35.1363 | -29.67
19.7351 | -35.1367 | -29.6763
19.7412 | -35.1372 | -29.6835
19.7473 | -35.1376 | -29.6901
19.7534 | -35.1381 | -29.6971
19.7595 | -35.1385 | -29.7044
19.7656 | -35.139 | -29.7112
19.7717 | -35.1394 | -29.7177
19.7778 | -35.1399 | -29.7247
19.7839 | -35.1404 | -29.7314
19.79 | -35.1408 | -29.7384
19.7961 | -35.1413 | -29.7454
19.8022 | -35.1418 | -29.7517
19.8083 | -35.1423 | -29.7582
19.8144 | -35.1428 | -29.765
19.8205 | -35.1432 | -29.772
19.8266 | -35.1437 | -29.7789
19.8327 | -35.1441 | -29.7858
19.8388 | -35.1445 | -29.7926
19.845 | -35.145 | -29.7995
19.8511 | -35.1454 | -29.8067
19.8572 | -35.1459 | -29.8134
19.8633 | -35.1463 | -29.8204
19.8694 | -35.1467 | -29.8274
19.8755 | -35.1471 | -29.834
19.8816 | -35.1476 | -29.8411
19.8877 | -35.148 | -29.8479
19.8938 | -35.1485 | -29.8548
19.8999 | -35.1489 | -29.8619
19.906 | -35.1494 | -29.8688
19.9121 | -35.1498 | -29.8754
19.9182 | -35.1502 | -29.8825
19.9243 | -35.1507 | -29.8894
19.9304 | -35.1511 | -29.8961
19.9365 | -35.1516 | -29.9032
19.9426 | -35.152 | -29.9099
19.9487 | -35.1524 | -29.9165
19.9548 | -35.1529 | -29.9236
19.9609 | -35.1533 | -29.9305
19.967 | -35.1538 | -29.9372
19.9731 | -35.1542 | -29.9439
19.9792 | -35.1547 | -29.9505
19.9853 | -35.1551 | -29.9573
19.9915 | -35.1556 | -29.9639
19.9976 | -35.156 | -29.9709
20.0037 | -35.1565 | -29.9776
20.0098 | -35.157 | -29.9841
20.0159 | -35.1574 | -29.9907
20.022 | -35.1578 | -29.9971
20.0281 | -35.1583 | -30.004
20.0342 | -35.1588 | -30.0107
20.0403 | -35.1592 | -30.0175
20.0464 | -35.1596 | -30.0242
20.0525 | -35.16 | -30.0307
20.0586 | -35.1604 | -30.0375
20.0647 | -35.1608 | -30.0446
20.0708 | -35.1613 | -30.0515
20.0769 | -35.1618 | -30.0587
20.083 | -35.1622 | -30.0652
20.0891 | -35.1626 | -30.072
20.0952 | -35.163 | -30.0788
20.1013 | -35.1635 | -30.0855
20.1074 | -35.164 | -30.092
20.1135 | -35.1645 | -30.0989
20.1196 | -35.1649 | -30.1053
20.1257 | -35.1653 | -30.112
20.1319 | -35.1658 | -30.1185
20.138 | -35.1662 | -30.1254
20.1441 | -35.1666 | -30.1318
20.1502 | -35.1671 | -30.1386
20.1563 | -35.1675 | -30.1458
20.1624 | -35.1679 | -30.1524
20.1685 | -35.1684 | -30.1593
20.1746 | -35.1688 | -30.1664
20.1807 | -35.1692 | -30.1731
20.1868 | -35.1697 | -30.1803
20.1929 | -35.1701 | -30.1865
20.199 | -35.1706 | -30.1935
20.2051 | -35.171 | -30.2001
20.2112 | -35.1715 | -30.2069
20.2173 | -35.1719 | -30.2135
20.2234 | -35.1724 | -30.2201
20.2295 | -35.1728 | -30.2267
20.2356 | -35.1733 | -30.2333
20.2417 | -35.1737 | -30.2403
20.2478 | -35.1741 | -30.2469
20.2539 | -35.1745 | -30.2539
20.26 | -35.1749 | -30.2608
20.2661 | -35.1754 | -30.2674
20.2723 | -35.1759 | -30.2741
20.2784 | -35.1763 | -30.2808
20.2845 | -35.1768 | -30.2876
20.2906 | -35.1772 | -30.2941
20.2967 | -35.1776 | -30.301
20.3028 | -35.1781 | -30.3076
20.3089 | -35.1785 | -30.3143
20.315 | -35.179 | -30.321
20.3211 | -35.1794 | -30.3276
20.3272 | -35.1798 | -30.3344
20.3333 | -35.1803 | -30.3407
20.3394 | -35.1807 | -30.3474
20.3455 | -35.1812 | -30.3543
20.3516 | -35.1816 | -30.3607
20.3577 | -35.1821 | -30.3675
20.3638 | -35.1825 | -30.3744
20.3699 | -35.1829 | -30.3804
20.376 | -35.1834 | -30.3869
20.3821 | -35.1839 | -30.3934
20.3882 | -35.1843 | -30.3999
20.3943 | -35.1848 | -30.4065
20.4004 | -35.1852 | -30.4127
20.4065 | -35.1856 | -30.4195
20.4126 | -35.1861 | -30.4259
20.4188 | -35.1865 | -30.4326
20.4249 | -35.1869 | -30.4389
20.431 | -35.1873 | -30.4453
20.4371 | -35.1877 | -30.4516
20.4432 | -35.1881 | -30.4585
20.4493 | -35.1886 | -30.4652
20.4554 | -35.1891 | -30.4717
20.4615 | -35.1895 | -30.4787
20.4676 | -35.1899 | -30.4854
20.4737 | -35.1904 | -30.492
20.4798 | -35.1908 | -30.4987
20.4859 | -35.1912 | -30.5054
20.492 | -35.1916 | -30.512
20.4981 | -35.192 | -30.5186
20.5042 | -35.1924 | -30.5252
20.5103 | -35.1929 | -30.5318
20.5164 | -35.1933 | -30.5385
20.5225 | -35.1938 | -30.5449
20.5286 | -35.1942 | -30.5517
20.5347 | -35.1947 | -30.558
20.5408 | -35.1951 | -30.5644
20.5469 | -35.1955 | -30.5711
20.553 | -35.1959 | -30.5776
20.5592 | -35.1964 | -30.5842
20.5653 | -35.1968 | -30.5907
20.5714 | -35.1972 | -30.5974
20.5775 | -35.1976 | -30.6031
20.5836 | -35.198 | -30.6094
20.5897 | -35.1985 | -30.6161
20.5958 | -35.1989 | -30.6228
20.6019 | -35.1993 | -30.6296
20.608 | -35.1998 | -30.6365
20.6141 | -35.2002 | -30.6425
20.6202 | -35.2007 | -30.649
20.6263 | -35.2011 | -30.6557
20.6324 | -35.2016 | -30.6629
20.6385 | -35.202 | -30.6695
20.6446 | -35.2024 | -30.6759
20.6507 | -35.2028 | -30.6822
20.6568 | -35.2033 | -30.6888
20.6629 | -35.2037 | -30.6957
20.669 | -35.2041 | -30.7021
20.6751 | -35.2045 | -30.7085
20.6812 | -35.2049 | -30.7149
20.6873 | -35.2054 | -30.7214
20.6934 | -35.2057 | -30.7276
20.6995 | -35.2062 | -30.7347
20.7057 | -35.2067 | -30.7412
20.7118 | -35.2072 | -30.748
20.7179 | -35.2077 | -30.7546
20.724 | -35.208 | -30.761
20.7301 | -35.2085 | -30.7678
20.7362 | -35.2089 | -30.7744
20.7423 | -35.2093 | -30.7811
20.7484 | -35.2098 | -30.7876
20.7545 | -35.2103 | -30.7938
20.7606 | -35.2107 | -30.8002
20.7667 | -35.2111 | -30.8065
20.7728 | -35.2115 | -30.8129
20.7789 | -35.2119 | -30.8189
20.785 | -35.2124 | -30.8251
20.7911 | -35.2127 | -30.8315
20.7972 | -35.2132 | -30.838
20.8033 | -35.2136 | -30.8446
20.8094 | -35.214 | -30.8509
20.8155 | -35.2145 | -30.8576
20.8216 | -35.2149 | -30.8641
20.8277 | -35.2155 | -30.8708
20.8338 | -35.2159 | -30.8774
20.8399 | -35.2163 | -30.8839
20.8461 | -35.2168 | -30.8908
20.8522 | -35.2172 | -30.8975
20.8583 | -35.2176 | -30.9041
20.8644 | -35.218 | -30.9111
20.8705 | -35.2184 | -30.9179
20.8766 | -35.2188 | -30.9246
20.8827 | -35.2193 | -30.9314
20.8888 | -35.2198 | -30.938
20.8949 | -35.2203 | -30.9442
20.901 | -35.2206 | -30.9508
20.9071 | -35.2211 | -30.9577
20.9132 | -35.2216 | -30.9646
20.9193 | -35.222 | -30.9708
20.9254 | -35.2224 | -30.9773
20.9315 | -35.2229 | -30.9847
20.9376 | -35.2233 | -30.9911
20.9437 | -35.2237 | -30.9977
20.9498 | -35.2241 | -31.0043
20.9559 | -35.2246 | -31.011
20.962 | -35.225 | -31.0174
20.9681 | -35.2255 | -31.0237
20.9742 | -35.226 | -31.0301
20.9803 | -35.2264 | -31.0366
20.9864 | -35.2268 | -31.0433
20.9926 | -35.2272 | -31.0501
20.9987 | -35.2276 | -31.0567
21.0048 | -35.2281 | -31.0637
21.0109 | -35.2285 | -31.0701
21.017 | -35.229 | -31.0767
21.0231 | -35.2295 | -31.0833
21.0292 | -35.2299 | -31.0896
21.0353 | -35.2303 | -31.0959
21.0414 | -35.2307 | -31.1026
21.0475 | -35.2311 | -31.1091
21.0536 | -35.2315 | -31.1154
21.0597 | -35.2319 | -31.1216
21.0658 | -35.2324 | -31.1281
21.0719 | -35.2328 | -31.1341
21.078 | -35.2333 | -31.1407
21.0841 | -35.2338 | -31.1473
21.0902 | -35.2342 | -31.1537
21.0963 | -35.2346 | -31.1598
21.1024 | -35.235 | -31.1665
21.1085 | -35.2356 | -31.1734
21.1146 | -35.236 | -31.1794
21.1207 | -35.2365 | -31.186
21.1268 | -35.2368 | -31.1928
21.133 | -35.2373 | -31.1994
21.1391 | -35.2377 | -31.2058
21.1452 | -35.2381 | -31.2121
21.1513 | -35.2385 | -31.2182
21.1574 | -35.2389 | -31.2251
21.1635 | -35.2394 | -31.2316
21.1696 | -35.2398 | -31.2379
21.1757 | -35.2401 | -31.2441
21.1818 | -35.2406 | -31.2502
21.1879 | -35.241 | -31.2566
21.194 | -35.2414 | -31.2629
21.2001 | -35.2418 | -31.2693
21.2062 | -35.2423 | -31.276
21.2123 | -35.2428 | -31.2828
21.2184 | -35.2433 | -31.2892
21.2245 | -35.2437 | -31.2949
21.2306 | -35.2442 | -31.3016
21.2367 | -35.2446 | -31.3083
21.2428 | -35.245 | -31.3159
21.2489 | -35.2455 | -31.3216
21.255 | -35.2459 | -31.3281
21.2611 | -35.2463 | -31.3345
21.2672 | -35.2467 | -31.3409
21.2733 | -35.2472 | -31.3471
21.2795 | -35.2476 | -31.3532
21.2856 | -35.248 | -31.3593
21.2917 | -35.2485 | -31.3658
21.2978 | -35.2489 | -31.3723
21.3039 | -35.2493 | -31.3785
21.31 | -35.2497 | -31.3852
21.3161 | -35.2502 | -31.3921
21.3222 | -35.2506 | -31.3988
21.3283 | -35.2511 | -31.4051
21.3344 | -35.2515 | -31.4115
21.3405 | -35.252 | -31.4175
21.3466 | -35.2524 | -31.4236
21.3527 | -35.2528 | -31.4304
21.3588 | -35.2532 | -31.4368
21.3649 | -35.2537 | -31.4434
21.371 | -35.2541 | -31.4499
21.3771 | -35.2545 | -31.4562
21.3832 | -35.255 | -31.4622
21.3893 | -35.2555 | -31.4686
21.3954 | -35.2559 | -31.4748
21.4015 | -35.2563 | -31.4811
21.4076 | -35.2566 | -31.4871
21.4137 | -35.2571 | -31.4934
21.4199 | -35.2576 | -31.4998
21.426 | -35.258 | -31.5063
21.4321 | -35.2584 | -31.5129
21.4382 | -35.2588 | -31.5192
21.4443 | -35.2592 | -31.5254
21.4504 | -35.2596 | -31.5315
21.4565 | -35.2601 | -31.5378
21.4626 | -35.2605 | -31.5441
21.4687 | -35.2609 | -31.5507
21.4748 | -35.2614 | -31.5569
21.4809 | -35.2619 | -31.5629
21.487 | -35.2622 | -31.5695
21.4931 | -35.2626 | -31.5755
21.4992 | -35.263 | -31.582
21.5053 | -35.2634 | -31.5883
21.5114 | -35.2639 | -31.5944
21.5175 | -35.2643 | -31.6007
21.5236 | -35.2647 | -31.6066
21.5297 | -35.2651 | -31.6129
21.5358 | -35.2655 | -31.6194
21.5419 | -35.2659 | -31.6259
21.548 | -35.2663 | -31.6325
21.5541 | -35.2668 | -31.6391
21.5602 | -35.2672 | -31.6451
21.5664 | -35.2677 | -31.6513
21.5725 | -35.2681 | -31.6578
21.5786 | -35.2685 | -31.6645
21.5847 | -35.2689 | -31.6717
21.5908 | -35.2694 | -31.6781
21.5969 | -35.2698 | -31.6844
21.603 | -35.2702 | -31.6908
21.6091 | -35.2706 | -31.6973
21.6152 | -35.271 | -31.7039
21.6213 | -35.2714 | -31.7103
21.6274 | -35.2718 | -31.717
21.6335 | -35.2722 | -31.7233
21.6396 | -35.2727 | -31.7295
21.6457 | -35.2731 | -31.736
21.6518 | -35.2735 | -31.7423
21.6579 | -35.2739 | -31.7483
21.664 | -35.2743 | -31.7549
21.6701 | -35.2746 | -31.7609
21.6762 | -35.2751 | -31.767
21.6823 | -35.2755 | -31.7735
21.6884 | -35.2759 | -31.7802
21.6945 | -35.2763 | -31.7869
21.7006 | -35.2767 | -31.7936
21.7068 | -35.277 | -31.8001
21.7129 | -35.2774 | -31.8063
21.719 | -35.2779 | -31.8123
21.7251 | -35.2782 | -31.8186
21.7312 | -35.2786 | -31.8247
21.7373 | -35.279 | -31.831
21.7434 | -35.2795 | -31.8372
21.7495 | -35.2799 | -31.8436
21.7556 | -35.2804 | -31.8505
21.7617 | -35.2808 | -31.8571
21.7678 | -35.2812 | -31.8632
21.7739 | -35.2817 | -31.8699
21.78 | -35.2821 | -31.8765
21.7861 | -35.2825 | -31.883
21.7922 | -35.283 | -31.8895
21.7983 | -35.2834 | -31.8959
21.8044 | -35.2838 | -31.902
21.8105 | -35.2843 | -31.9084
21.8166 | -35.2847 | -31.9151
21.8227 | -35.2851 | -31.9214
21.8288 | -35.2855 | -31.928
21.8349 | -35.286 | -31.9346
21.841 | -35.2864 | -31.9408
21.8471 | -35.2868 | -31.9474
21.8533 | -35.2872 | -31.9536
21.8594 | -35.2877 | -31.9598
21.8655 | -35.2881 | -31.9657
21.8716 | -35.2884 | -31.9715
21.8777 | -35.2888 | -31.9778
21.8838 | -35.2893 | -31.9842
21.8899 | -35.2897 | -31.9908
21.896 | -35.2901 | -31.997
21.9021 | -35.2905 | -32.0042
21.9082 | -35.291 | -32.0099
21.9143 | -35.2914 | -32.0172
21.9204 | -35.2918 | -32.0236
21.9265 | -35.2922 | -32.0292
21.9326 | -35.2926 | -32.0354
21.9387 | -35.293 | -32.0417
21.9448 | -35.2935 | -32.0485
21.9509 | -35.2939 | -32.0548
21.957 | -35.2944 | -32.0605
21.9631 | -35.2948 | -32.0666
21.9692 | -35.2952 | -32.0734
21.9753 | -35.2957 | -32.0803
21.9814 | -35.2961 | -32.087
21.9875 | -35.2965 | -32.0935
21.9937 | -35.2968 | -32.0995
21.9998 | -35.2973 | -32.1064
22.0059 | -35.2977 | -32.1129
22.012 | -35.2981 | -32.1193
22.0181 | -35.2985 | -32.1261
22.0242 | -35.2989 | -32.1324
22.0303 | -35.2994 | -32.1392
22.0364 | -35.2999 | -32.1453
22.0425 | -35.3003 | -32.152
22.0486 | -35.3007 | -32.1585
22.0547 | -35.3011 | -32.165
22.0608 | -35.3015 | -32.1716
22.0669 | -35.3019 | -32.1774
22.073 | -35.3024 | -32.1839
22.0791 | -35.3028 | -32.1899
22.0852 | -35.3033 | -32.1963
22.0913 | -35.3037 | -32.2031
22.0974 | -35.3041 | -32.209
22.1035 | -35.3045 | -32.2153
22.1096 | -35.3049 | -32.2215
22.1157 | -35.3053 | -32.2279
22.1218 | -35.3058 | -32.234
22.1279 | -35.3062 | -32.2397
22.134 | -35.3066 | -32.2458
22.1402 | -35.307 | -32.2517
22.1463 | -35.3074 | -32.2583
22.1524 | -35.3079 | -32.2645
22.1585 | -35.3084 | -32.2707
22.1646 | -35.3087 | -32.277
22.1707 | -35.3091 | -32.2834
22.1768 | -35.3095 | -32.2894
22.1829 | -35.31 | -32.2956
22.189 | -35.3104 | -32.3022
22.1951 | -35.3109 | -32.3089
22.2012 | -35.3113 | -32.3156
22.2073 | -35.3117 | -32.3213
22.2134 | -35.3121 | -32.3271
22.2195 | -35.3126 | -32.3333
22.2256 | -35.3131 | -32.3391
22.2317 | -35.3134 | -32.3451
22.2378 | -35.3138 | -32.3517
22.2439 | -35.3142 | -32.3575
22.25 | -35.3146 | -32.3637
22.2561 | -35.315 | -32.3703
22.2622 | -35.3154 | -32.3767
22.2683 | -35.3158 | -32.3829
22.2744 | -35.3163 | -32.3892
22.2806 | -35.3167 | -32.3955
22.2867 | -35.3171 | -32.402
22.2928 | -35.3176 | -32.4084
22.2989 | -35.318 | -32.4139
22.305 | -35.3185 | -32.4203
22.3111 | -35.3189 | -32.427
22.3172 | -35.3193 | -32.4332
22.3233 | -35.3197 | -32.4396
22.3294 | -35.3202 | -32.4456
22.3355 | -35.3206 | -32.4519
22.3416 | -35.321 | -32.4579
22.3477 | -35.3214 | -32.4642
22.3538 | -35.3218 | -32.4704
22.3599 | -35.3221 | -32.4766
22.366 | -35.3226 | -32.4829
22.3721 | -35.323 | -32.4891
22.3782 | -35.3235 | -32.4954
22.3843 | -35.3238 | -32.5016
22.3904 | -35.3242 | -32.5083
22.3965 | -35.3247 | -32.5149
22.4026 | -35.3252 | -32.5214
22.4087 | -35.3256 | -32.5277
22.4148 | -35.326 | -32.5337
22.4209 | -35.3264 | -32.5399
22.4271 | -35.3268 | -32.5462
22.4332 | -35.3272 | -32.5524
22.4393 | -35.3276 | -32.5582
22.4454 | -35.328 | -32.565
22.4515 | -35.3284 | -32.5713
22.4576 | -35.3288 | -32.5777
22.4637 | -35.3292 | -32.5838
22.4698 | -35.3296 | -32.5903
22.4759 | -35.33 | -32.5969
22.482 | -35.3304 | -32.6033
22.4881 | -35.3309 | -32.6097
22.4942 | -35.3313 | -32.6157
22.5003 | -35.3316 | -32.6221
22.5064 | -35.3321 | -32.6286
22.5125 | -35.3325 | -32.6344
22.5186 | -35.3329 | -32.6408
22.5247 | -35.3332 | -32.6472
22.5308 | -35.3336 | -32.6532
22.5369 | -35.334 | -32.6599
22.543 | -35.3344 | -32.666
22.5491 | -35.3348 | -32.6721
22.5552 | -35.3353 | -32.6783
22.5613 | -35.3357 | -32.6849
22.5675 | -35.3361 | -32.6912
22.5736 | -35.3365 | -32.6979
22.5797 | -35.3369 | -32.7038
22.5858 | -35.3372 | -32.7102
22.5919 | -35.3376 | -32.7169
22.598 | -35.338 | -32.7234
22.6041 | -35.3384 | -32.73
22.6102 | -35.3388 | -32.7364
22.6163 | -35.3392 | -32.7425
22.6224 | -35.3396 | -32.7487
22.6285 | -35.3401 | -32.7548
22.6346 | -35.3405 | -32.7612
22.6407 | -35.3409 | -32.7673
22.6468 | -35.3414 | -32.7737
22.6529 | -35.3418 | -32.7801
22.659 | -35.3422 | -32.7868
22.6651 | -35.3427 | -32.7931
22.6712 | -35.3431 | -32.7996
22.6773 | -35.3436 | -32.8059
22.6834 | -35.344 | -32.8122
22.6895 | -35.3444 | -32.8179
22.6956 | -35.3448 | -32.8243
22.7017 | -35.3452 | -32.8305
22.7079 | -35.3456 | -32.8366
22.714 | -35.346 | -32.8429
22.7201 | -35.3464 | -32.8489
22.7262 | -35.3468 | -32.8546
22.7323 | -35.3472 | -32.8612
22.7384 | -35.3477 | -32.8674
22.7445 | -35.348 | -32.8736
22.7506 | -35.3484 | -32.8795
22.7567 | -35.3488 | -32.8859
22.7628 | -35.3493 | -32.8913
22.7689 | -35.3498 | -32.8975
22.775 | -35.3502 | -32.9037
22.7811 | -35.3506 | -32.9105
22.7872 | -35.351 | -32.9168
22.7933 | -35.3513 | -32.9229
22.7994 | -35.3517 | -32.9289
22.8055 | -35.3522 | -32.9353
22.8116 | -35.3526 | -32.9415
22.8177 | -35.353 | -32.9479
22.8238 | -35.3534 | -32.9538
22.8299 | -35.3539 | -32.9601
22.836 | -35.3543 | -32.9661
22.8421 | -35.3547 | -32.9724
22.8482 | -35.355 | -32.9784
22.8544 | -35.3554 | -32.9846
22.8605 | -35.3558 | -32.9906
22.8666 | -35.3561 | -32.9968
22.8727 | -35.3565 | -33.0033
22.8788 | -35.3569 | -33.0087
22.8849 | -35.3574 | -33.0146
22.891 | -35.3578 | -33.0211
22.8971 | -35.3581 | -33.0274
22.9032 | -35.3587 | -33.0333
22.9093 | -35.3591 | -33.0397
22.9154 | -35.3595 | -33.046
22.9215 | -35.3599 | -33.0519
22.9276 | -35.3603 | -33.058
22.9337 | -35.3608 | -33.0641
22.9398 | -35.3612 | -33.0707
22.9459 | -35.3616 | -33.0771
22.952 | -35.3621 | -33.0832
22.9581 | -35.3625 | -33.0898
22.9642 | -35.3628 | -33.0961
22.9703 | -35.3633 | -33.102
22.9764 | -35.3638 | -33.1081
22.9825 | -35.3642 | -33.1144
22.9886 | -35.3646 | -33.1201
22.9948 | -35.365 | -33.126
23.0009 | -35.3654 | -33.1319
23.007 | -35.3658 | -33.1385
23.0131 | -35.3661 | -33.1445
23.0192 | -35.3665 | -33.1502
23.0253 | -35.367 | -33.1565
23.0314 | -35.3673 | -33.1627
23.0375 | -35.3677 | -33.1693
23.0436 | -35.3681 | -33.175
23.0497 | -35.3685 | -33.181
23.0558 | -35.3689 | -33.1872
23.0619 | -35.3693 | -33.1933
23.068 | -35.3697 | -33.1991
23.0741 | -35.3701 | -33.2049
23.0802 | -35.3705 | -33.2107
23.0863 | -35.371 | -33.2171
23.0924 | -35.3714 | -33.2231
23.0985 | -35.3717 | -33.2293
23.1046 | -35.3721 | -33.2353
23.1107 | -35.3725 | -33.2412
23.1168 | -35.3729 | -33.2476
23.1229 | -35.3733 | -33.2538
23.129 | -35.3737 | -33.2596
23.1351 | -35.3741 | -33.2659
23.1413 | -35.3745 | -33.2721
23.1474 | -35.3749 | -33.2777
23.1535 | -35.3753 | -33.2839
23.1596 | -35.3757 | -33.2902
23.1657 | -35.3761 | -33.2966
23.1718 | -35.3765 | -33.303
23.1779 | -35.3769 | -33.3095
23.184 | -35.3773 | -33.3155
23.1901 | -35.3777 | -33.3217
23.1962 | -35.3781 | -33.3277
23.2023 | -35.3785 | -33.3338
23.2084 | -35.3789 | -33.34
23.2145 | -35.3793 | -33.3455
23.2206 | -35.3796 | -33.3519
23.2267 | -35.3801 | -33.3579
23.2328 | -35.3804 | -33.3639
23.2389 | -35.3808 | -33.3697
23.245 | -35.3812 | -33.376
23.2511 | -35.3815 | -33.3818
23.2572 | -35.3819 | -33.3879
23.2633 | -35.3823 | -33.3939
23.2694 | -35.3827 | -33.3999
23.2755 | -35.3831 | -33.4062
23.2817 | -35.3835 | -33.4125
23.2878 | -35.3839 | -33.418
23.2939 | -35.3844 | -33.4243
23.3 | -35.3847 | -33.4304
23.3061 | -35.3852 | -33.4364
23.3122 | -35.3856 | -33.4427
23.3183 | -35.386 | -33.4487
23.3244 | -35.3864 | -33.4547
23.3305 | -35.3868 | -33.4605
23.3366 | -35.3871 | -33.4669
23.3427 | -35.3875 | -33.4732
23.3488 | -35.3878 | -33.4792
23.3549 | -35.3883 | -33.4856
23.361 | -35.3886 | -33.4914
23.3671 | -35.389 | -33.4977
23.3732 | -35.3894 | -33.5042
23.3793 | -35.3898 | -33.5104
23.3854 | -35.3902 | -33.5164
23.3915 | -35.3907 | -33.5221
23.3976 | -35.3911 | -33.5285
23.4037 | -35.3915 | -33.5347
23.4098 | -35.392 | -33.541
23.4159 | -35.3924 | -33.5473
23.422 | -35.3929 | -33.5534
23.4282 | -35.3933 | -33.5593
23.4343 | -35.3937 | -33.5653
23.4404 | -35.3941 | -33.5715
23.4465 | -35.3944 | -33.5773
23.4526 | -35.3948 | -33.5834
23.4587 | -35.3952 | -33.5894
23.4648 | -35.3956 | -33.5953
23.4709 | -35.396 | -33.601
23.477 | -35.3965 | -33.6071
23.4831 | -35.3968 | -33.613
23.4892 | -35.3972 | -33.6189
23.4953 | -35.3976 | -33.6248
23.5014 | -35.398 | -33.6305
23.5075 | -35.3984 | -33.6364
23.5136 | -35.3988 | -33.6422
23.5197 | -35.3992 | -33.6479
23.5258 | -35.3996 | -33.654
23.5319 | -35.3999 | -33.6599
23.538 | -35.4003 | -33.6658
23.5441 | -35.4008 | -33.6718
23.5502 | -35.4012 | -33.6779
23.5563 | -35.4016 | -33.6838
23.5624 | -35.4019 | -33.6899
23.5686 | -35.4024 | -33.6957
23.5747 | -35.4028 | -33.7014
23.5808 | -35.4032 | -33.7076
23.5869 | -35.4036 | -33.7135
23.593 | -35.404 | -33.7192
23.5991 | -35.4045 | -33.7251
23.6052 | -35.4049 | -33.7312
23.6113 | -35.4053 | -33.7372
23.6174 | -35.4056 | -33.7427
23.6235 | -35.4061 | -33.7486
23.6296 | -35.4065 | -33.7546
23.6357 | -35.4069 | -33.7604
23.6418 | -35.4072 | -33.7659
23.6479 | -35.4076 | -33.7716
23.654 | -35.408 | -33.7774
23.6601 | -35.4084 | -33.7834
23.6662 | -35.4087 | -33.7891
23.6723 | -35.409 | -33.7951
23.6784 | -35.4094 | -33.801
23.6845 | -35.4098 | -33.8068
23.6906 | -35.4102 | -33.8129
23.6967 | -35.4107 | -33.8188
23.7028 | -35.4111 | -33.8251
23.7089 | -35.4115 | -33.8304
23.7151 | -35.4119 | -33.8359
23.7212 | -35.4123 | -33.8418
23.7273 | -35.4127 | -33.848
23.7334 | -35.4131 | -33.8541
23.7395 | -35.4135 | -33.8601
23.7456 | -35.4138 | -33.8657
23.7517 | -35.4142 | -33.8719
23.7578 | -35.4145 | -33.8778
23.7639 | -35.4148 | -33.8843
23.77 | -35.4152 | -33.8902
23.7761 | -35.4155 | -33.8961
23.7822 | -35.4159 | -33.9018
23.7883 | -35.4164 | -33.9079
23.7944 | -35.4168 | -33.914
23.8005 | -35.4171 | -33.9201
23.8066 | -35.4175 | -33.9261
23.8127 | -35.4179 | -33.9322
23.8188 | -35.4183 | -33.9383
23.8249 | -35.4187 | -33.9443
23.831 | -35.419 | -33.9499
23.8371 | -35.4194 | -33.9562
23.8432 | -35.4198 | -33.9621
23.8493 | -35.4201 | -33.9679
23.8555 | -35.4205 | -33.9739
23.8616 | -35.4209 | -33.9801
23.8677 | -35.4213 | -33.9861
23.8738 | -35.4217 | -33.9918
23.8799 | -35.4221 | -33.9979
23.886 | -35.4225 | -34.003
23.8921 | -35.4228 | -34.0088
23.8982 | -35.4232 | -34.0146
23.9043 | -35.4235 | -34.0204
23.9104 | -35.424 | -34.0262
23.9165 | -35.4244 | -34.0326
23.9226 | -35.4247 | -34.0384
23.9287 | -35.4251 | -34.0441
23.9348 | -35.4255 | -34.0503
23.9409 | -35.4258 | -34.0563
23.947 | -35.4263 | -34.0622
23.9531 | -35.4267 | -34.0674
23.9592 | -35.4271 | -34.0738
23.9653 | -35.4275 | -34.0797
23.9714 | -35.4278 | -34.0855
23.9775 | -35.4282 | -34.0917
23.9836 | -35.4286 | -34.0975
23.9897 | -35.429 | -34.1029
23.9958 | -35.4293 | -34.1085
24.002 | -35.4297 | -34.1141
24.0081 | -35.4301 | -34.1197
24.0142 | -35.4305 | -34.1258
24.0203 | -35.4309 | -34.1316
24.0264 | -35.4313 | -34.1373
24.0325 | -35.4317 | -34.1429
24.0386 | -35.4321 | -34.1492
24.0447 | -35.4324 | -34.1547
24.0508 | -35.4328 | -34.1607
24.0569 | -35.4332 | -34.1663
24.063 | -35.4336 | -34.1714
24.0691 | -35.434 | -34.1774
24.0752 | -35.4344 | -34.1832
24.0813 | -35.4348 | -34.1892
24.0874 | -35.4351 | -34.195
24.0935 | -35.4355 | -34.201
24.0996 | -35.4358 | -34.2065
24.1057 | -35.4362 | -34.2123
24.1118 | -35.4365 | -34.2185
24.1179 | -35.4368 | -34.224
24.124 | -35.4372 | -34.2299
24.1301 | -35.4375 | -34.235
24.1362 | -35.4379 | -34.241
24.1424 | -35.4383 | -34.2472
24.1485 | -35.4387 | -34.2526
24.1546 | -35.439 | -34.2584
24.1607 | -35.4394 | -34.2643
24.1668 | -35.4397 | -34.2697
24.1729 | -35.4401 | -34.2758
24.179 | -35.4405 | -34.2817
24.1851 | -35.4408 | -34.2875
24.1912 | -35.4412 | -34.2934
24.1973 | -35.4416 | -34.2991
24.2034 | -35.442 | -34.3041
24.2095 | -35.4424 | -34.31
24.2156 | -35.4428 | -34.3157
24.2217 | -35.4432 | -34.3215
24.2278 | -35.4436 | -34.3272
24.2339 | -35.4439 | -34.3326
24.24 | -35.4443 | -34.3378
24.2461 | -35.4446 | -34.3436
24.2522 | -35.445 | -34.3494
24.2583 | -35.4454 | -34.3548
24.2644 | -35.4458 | -34.3606
24.2705 | -35.4462 | -34.3664
24.2766 | -35.4466 | -34.372
24.2827 | -35.4469 | -34.3777
24.2889 | -35.4473 | -34.3833
24.295 | -35.4477 | -34.3889
24.3011 | -35.448 | -34.3948
24.3072 | -35.4483 | -34.4006
24.3133 | -35.4487 | -34.4066
24.3194 | -35.4491 | -34.4122
24.3255 | -35.4495 | -34.418
24.3316 | -35.4499 | -34.4233
24.3377 | -35.4502 | -34.4291
24.3438 | -35.4505 | -34.4346
24.3499 | -35.4509 | -34.4403
24.356 | -35.4512 | -34.4464
24.3621 | -35.4516 | -34.4521
24.3682 | -35.4519 | -34.4578
24.3743 | -35.4523 | -34.4631
24.3804 | -35.4527 | -34.4684
24.3865 | -35.453 | -34.4744
24.3926 | -35.4533 | -34.4802
24.3987 | -35.4537 | -34.4863
24.4048 | -35.454 | -34.4922
24.4109 | -35.4544 | -34.4973
24.417 | -35.4549 | -34.503
24.4231 | -35.4553 | -34.5086
24.4293 | -35.4556 | -34.5146
24.4354 | -35.456 | -34.5204
24.4415 | -35.4564 | -34.5262
24.4476 | -35.4567 | -34.5313
24.4537 | -35.4571 | -34.537
24.4598 | -35.4575 | -34.5428
24.4659 | -35.4579 | -34.5487
24.472 | -35.4583 | -34.5541
24.4781 | -35.4586 | -34.5595
24.4842 | -35.459 | -34.565
24.4903 | -35.4593 | -34.5706
24.4964 | -35.4598 | -34.5763
24.5025 | -35.4601 | -34.5821
24.5086 | -35.4605 | -34.5878
24.5147 | -35.4608 | -34.5934
24.5208 | -35.4612 | -34.5984
24.5269 | -35.4616 | -34.6036
24.533 | -35.462 | -34.6091
24.5391 | -35.4623 | -34.6147
24.5452 | -35.4627 | -34.6204
24.5513 | -35.463 | -34.6255
24.5574 | -35.4634 | -34.6311
24.5635 | -35.4637 | -34.6367
24.5696 | -35.4641 | -34.6425
24.5758 | -35.4645 | -34.6481
24.5819 | -35.4648 | -34.6538
24.588 | -35.4652 | -34.6592
24.5941 | -35.4656 | -34.6652
24.6002 | -35.466 | -34.6705
24.6063 | -35.4664 | -34.6763
24.6124 | -35.4667 | -34.6822
24.6185 | -35.4671 | -34.6879
24.6246 | -35.4674 | -34.6932
24.6307 | -35.4678 | -34.6993
24.6368 | -35.4681 | -34.7051
24.6429 | -35.4685 | -34.7107
24.649 | -35.4688 | -34.7164
24.6551 | -35.4692 | -34.7218
24.6612 | -35.4695 | -34.7272
24.6673 | -35.4699 | -34.7329
24.6734 | -35.4702 | -34.7383
24.6795 | -35.4705 | -34.7443
24.6856 | -35.4709 | -34.7497
24.6917 | -35.4713 | -34.7552
24.6978 | -35.4716 | -34.7605
24.7039 | -35.472 | -34.7662
24.71 | -35.4723 | -34.7717
24.7162 | -35.4727 | -34.7772
24.7223 | -35.4731 | -34.7825
24.7284 | -35.4734 | -34.7878
24.7345 | -35.4738 | -34.7936
24.7406 | -35.4742 | -34.7993
24.7467 | -35.4745 | -34.8045
24.7528 | -35.4748 | -34.8101
24.7589 | -35.4752 | -34.8153
24.765 | -35.4755 | -34.8208
24.7711 | -35.4758 | -34.8263
24.7772 | -35.4762 | -34.8321
24.7833 | -35.4765 | -34.8375
24.7894 | -35.4769 | -34.8429
24.7955 | -35.4773 | -34.8484
24.8016 | -35.4776 | -34.8536
24.8077 | -35.4779 | -34.8595
24.8138 | -35.4783 | -34.8651
24.8199 | -35.4787 | -34.8707
24.826 | -35.4791 | -34.8765
24.8321 | -35.4794 | -34.882
24.8382 | -35.4797 | -34.887
24.8443 | -35.4801 | -34.8924
24.8504 | -35.4805 | -34.8978
24.8565 | -35.4808 | -34.9035
24.8627 | -35.4811 | -34.9088
24.8688 | -35.4815 | -34.9142
24.8749 | -35.4818 | -34.9198
24.881 | -35.4822 | -34.9255
24.8871 | -35.4825 | -34.9311
24.8932 | -35.4829 | -34.9368
24.8993 | -35.4833 | -34.9425
24.9054 | -35.4836 | -34.9476
24.9115 | -35.4839 | -34.9527
24.9176 | -35.4842 | -34.9584
24.9237 | -35.4846 | -34.9638
24.9298 | -35.4849 | -34.9693
24.9359 | -35.4853 | -34.9751
24.942 | -35.4857 | -34.9799
24.9481 | -35.486 | -34.9852
24.9542 | -35.4863 | -34.991
24.9603 | -35.4866 | -34.9964
24.9664 | -35.487 | -35.0018
24.9725 | -35.4873 | -35.007
24.9786 | -35.4877 | -35.0121
24.9847 | -35.4881 | -35.0176
24.9908 | -35.4884 | -35.0229
24.9969 | -35.4888 | -35.0286
25.0031 | -35.4891 | -35.034
25.0092 | -35.4894 | -35.039
25.0153 | -35.4897 | -35.0442
25.0214 | -35.4901 | -35.05
25.0275 | -35.4904 | -35.0554
25.0336 | -35.4908 | -35.0608
25.0397 | -35.4912 | -35.0662
25.0458 | -35.4915 | -35.0712
25.0519 | -35.4919 | -35.0766
25.058 | -35.4922 | -35.0822
25.0641 | -35.4925 | -35.0874
25.0702 | -35.4929 | -35.0928
25.0763 | -35.4933 | -35.0981
25.0824 | -35.4936 | -35.1032
25.0885 | -35.4939 | -35.1085
25.0946 | -35.4943 | -35.1139
25.1007 | -35.4946 | -35.1194
25.1068 | -35.495 | -35.1243
25.1129 | -35.4953 | -35.1295
25.119 | -35.4956 | -35.1345
25.1251 | -35.496 | -35.1401
25.1312 | -35.4963 | -35.1458
25.1373 | -35.4967 | -35.1512
25.1435 | -35.497 | -35.1567
25.1496 | -35.4974 | -35.1615
25.1557 | -35.4977 | -35.1669
25.1618 | -35.4981 | -35.172
25.1679 | -35.4984 | -35.1777
25.174 | -35.4987 | -35.183
25.1801 | -35.4991 | -35.1885
25.1862 | -35.4994 | -35.1937
25.1923 | -35.4997 | -35.1991
25.1984 | -35.5001 | -35.2045
25.2045 | -35.5004 | -35.2099
25.2106 | -35.5008 | -35.2151
25.2167 | -35.5011 | -35.2208
25.2228 | -35.5014 | -35.2256
25.2289 | -35.5018 | -35.231
25.235 | -35.5021 | -35.2365
25.2411 | -35.5025 | -35.2419
25.2472 | -35.5028 | -35.2475
25.2533 | -35.5032 | -35.2527
25.2594 | -35.5035 | -35.2579
25.2655 | -35.5038 | -35.2632
25.2716 | -35.5041 | -35.2682
25.2777 | -35.5045 | -35.2736
25.2838 | -35.5049 | -35.2791
25.29 | -35.5052 | -35.2846
25.2961 | -35.5056 | -35.2897
25.3022 | -35.5059 | -35.295
25.3083 | -35.5062 | -35.3002
25.3144 | -35.5066 | -35.3056
25.3205 | -35.507 | -35.311
25.3266 | -35.5074 | -35.3163
25.3327 | -35.5077 | -35.3221
25.3388 | -35.508 | -35.3277
25.3449 | -35.5083 | -35.3328
25.351 | -35.5086 | -35.3382
25.3571 | -35.509 | -35.3437
25.3632 | -35.5093 | -35.3489
25.3693 | -35.5096 | -35.3544
25.3754 | -35.51 | -35.3597
25.3815 | -35.5104 | -35.365
25.3876 | -35.5107 | -35.3703
25.3937 | -35.5111 | -35.3754
25.3998 | -35.5114 | -35.3805
25.4059 | -35.5117 | -35.386
25.412 | -35.5121 | -35.3912
25.4181 | -35.5124 | -35.3963
25.4242 | -35.5128 | -35.4014
25.4304 | -35.5131 | -35.4069
25.4365 | -35.5134 | -35.4119
25.4426 | -35.5137 | -35.4172
25.4487 | -35.5141 | -35.4227
25.4548 | -35.5144 | -35.428
25.4609 | -35.5148 | -35.4333
25.467 | -35.5151 | -35.4384
25.4731 | -35.5155 | -35.444
25.4792 | -35.5158 | -35.449
25.4853 | -35.5162 | -35.4542
25.4914 | -35.5165 | -35.4598
25.4975 | -35.5168 | -35.4651
25.5036 | -35.5171 | -35.4704
25.5097 | -35.5175 | -35.4756
25.5158 | -35.5179 | -35.4807
25.5219 | -35.5182 | -35.4861
25.528 | -35.5185 | -35.4911
25.5341 | -35.5188 | -35.4965
25.5402 | -35.5191 | -35.502
25.5463 | -35.5195 | -35.5075
25.5524 | -35.5198 | -35.5125
25.5585 | -35.5202 | -35.5177
25.5646 | -35.5205 | -35.5228
25.5707 | -35.5208 | -35.528
25.5769 | -35.5211 | -35.533
25.583 | -35.5215 | -35.5382
25.5891 | -35.5218 | -35.5434
25.5952 | -35.5221 | -35.5484
25.6013 | -35.5224 | -35.5538
25.6074 | -35.5228 | -35.5585
25.6135 | -35.5231 | -35.5633
25.6196 | -35.5235 | -35.5689
25.6257 | -35.5238 | -35.5742
25.6318 | -35.5241 | -35.5797
25.6379 | -35.5244 | -35.5851
25.644 | -35.5248 | -35.5902
25.6501 | -35.5251 | -35.5955
25.6562 | -35.5255 | -35.6006
25.6623 | -35.5259 | -35.6058
25.6684 | -35.5262 | -35.611
25.6745 | -35.5266 | -35.6161
25.6806 | -35.5269 | -35.621
25.6867 | -35.5273 | -35.6261
25.6928 | -35.5276 | -35.6313
25.6989 | -35.5279 | -35.6363
25.705 | -35.5283 | -35.6418
25.7111 | -35.5286 | -35.6471
25.7173 | -35.529 | -35.6522
25.7234 | -35.5293 | -35.6573
25.7295 | -35.5296 | -35.6625
25.7356 | -35.5299 | -35.6678
25.7417 | -35.5303 | -35.6729
25.7478 | -35.5306 | -35.6779
25.7539 | -35.5309 | -35.6831
25.76 | -35.5313 | -35.6884
25.7661 | -35.5316 | -35.6935
25.7722 | -35.5319 | -35.6987
25.7783 | -35.5323 | -35.704
25.7844 | -35.5326 | -35.7093
25.7905 | -35.533 | -35.7145
25.7966 | -35.5333 | -35.7195
25.8027 | -35.5336 | -35.7246
25.8088 | -35.5339 | -35.7298
25.8149 | -35.5342 | -35.735
25.821 | -35.5345 | -35.7398
25.8271 | -35.5348 | -35.7451
25.8332 | -35.5351 | -35.7505
25.8393 | -35.5355 | -35.7559
25.8454 | -35.5357 | -35.7611
25.8515 | -35.5361 | -35.7662
25.8576 | -35.5364 | -35.7715
25.8638 | -35.5367 | -35.7769
25.8699 | -35.537 | -35.7821
25.876 | -35.5373 | -35.7871
25.8821 | -35.5376 | -35.7925
25.8882 | -35.538 | -35.7976
25.8943 | -35.5383 | -35.8031
25.9004 | -35.5387 | -35.8081
25.9065 | -35.539 | -35.8134
25.9126 | -35.5393 | -35.8184
25.9187 | -35.5397 | -35.8236
25.9248 | -35.54 | -35.8284
25.9309 | -35.5403 | -35.834
25.937 | -35.5407 | -35.8395
25.9431 | -35.541 | -35.8449
25.9492 | -35.5413 | -35.85
25.9553 | -35.5416 | -35.8553
25.9614 | -35.5419 | -35.8601
25.9675 | -35.5423 | -35.8655
25.9736 | -35.5427 | -35.8705
25.9797 | -35.543 | -35.8757
25.9858 | -35.5433 | -35.8807
25.9919 | -35.5436 | -35.8862
25.998 | -35.5439 | -35.8912
26.0042 | -35.5442 | -35.8962
26.0103 | -35.5445 | -35.9012
26.0164 | -35.5448 | -35.9064
26.0225 | -35.5452 | -35.9114
26.0286 | -35.5455 | -35.9166
26.0347 | -35.5458 | -35.9213
26.0408 | -35.5461 | -35.9264
26.0469 | -35.5464 | -35.9314
26.053 | -35.5468 | -35.9368
26.0591 | -35.5471 | -35.9417
26.0652 | -35.5474 | -35.9463
26.0713 | -35.5477 | -35.9513
26.0774 | -35.548 | -35.956
26.0835 | -35.5483 | -35.9611
26.0896 | -35.5487 | -35.966
26.0957 | -35.549 | -35.9715
26.1018 | -35.5493 | -35.9763
26.1079 | -35.5496 | -35.9815
26.114 | -35.55 | -35.9867
26.1201 | -35.5503 | -35.9919
26.1262 | -35.5506 | -35.9969
26.1323 | -35.5509 | -36.0018
26.1384 | -35.5513 | -36.0065
26.1445 | -35.5516 | -36.0117
26.1507 | -35.5519 | -36.0169
26.1568 | -35.5522 | -36.022
26.1629 | -35.5526 | -36.0271
26.169 | -35.5529 | -36.032
26.1751 | -35.5532 | -36.0368
26.1812 | -35.5535 | -36.0418
26.1873 | -35.5539 | -36.047
26.1934 | -35.5542 | -36.0521
26.1995 | -35.5545 | -36.0571
26.2056 | -35.5548 | -36.0621
26.2117 | -35.5551 | -36.0674
26.2178 | -35.5554 | -36.0725
26.2239 | -35.5557 | -36.0777
26.23 | -35.556 | -36.0829
26.2361 | -35.5563 | -36.0882
26.2422 | -35.5566 | -36.0931
26.2483 | -35.5569 | -36.0982
26.2544 | -35.5572 | -36.1031
26.2605 | -35.5576 | -36.1081
26.2666 | -35.5579 | -36.1129
26.2727 | -35.5582 | -36.118
26.2788 | -35.5586 | -36.123
26.2849 | -35.5589 | -36.1281
26.2911 | -35.5592 | -36.1329
26.2972 | -35.5595 | -36.1378
26.3033 | -35.5599 | -36.1431
26.3094 | -35.5602 | -36.1483
26.3155 | -35.5605 | -36.1531
26.3216 | -35.5609 | -36.1583
26.3277 | -35.5612 | -36.1635
26.3338 | -35.5615 | -36.1684
26.3399 | -35.5618 | -36.1733
26.346 | -35.5622 | -36.178
26.3521 | -35.5625 | -36.1832
26.3582 | -35.5628 | -36.1882
26.3643 | -35.5631 | -36.1931
26.3704 | -35.5634 | -36.1983
26.3765 | -35.5637 | -36.2036
26.3826 | -35.564 | -36.2083
26.3887 | -35.5643 | -36.2135
26.3948 | -35.5647 | -36.2184
26.4009 | -35.565 | -36.2235
26.407 | -35.5653 | -36.2286
26.4131 | -35.5657 | -36.234
26.4192 | -35.566 | -36.2385
26.4253 | -35.5663 | -36.2439
26.4314 | -35.5666 | -36.2489
26.4376 | -35.5669 | -36.254
26.4437 | -35.5672 | -36.2593
26.4498 | -35.5675 | -36.2648
26.4559 | -35.5678 | -36.2693
26.462 | -35.5681 | -36.2743
26.4681 | -35.5683 | -36.2792
26.4742 | -35.5687 | -36.2846
26.4803 | -35.569 | -36.2895
26.4864 | -35.5693 | -36.2945
26.4925 | -35.5696 | -36.2993
26.4986 | -35.5699 | -36.3044
26.5047 | -35.5702 | -36.3095
26.5108 | -35.5705 | -36.3144
26.5169 | -35.5708 | -36.3193
26.523 | -35.5711 | -36.3242
26.5291 | -35.5715 | -36.3294
26.5352 | -35.5718 | -36.3342
26.5413 | -35.5721 | -36.3393
26.5474 | -35.5724 | -36.3447
26.5535 | -35.5727 | -36.35
26.5596 | -35.573 | -36.3548
26.5657 | -35.5733 | -36.36
26.5718 | -35.5737 | -36.3648
26.578 | -35.574 | -36.3697
26.5841 | -35.5743 | -36.3745
26.5902 | -35.5746 | -36.3794
26.5963 | -35.5748 | -36.384
26.6024 | -35.5752 | -36.389
26.6085 | -35.5755 | -36.3944
26.6146 | -35.5758 | -36.3996
26.6207 | -35.5761 | -36.4047
26.6268 | -35.5764 | -36.4099
26.6329 | -35.5767 | -36.4148
26.639 | -35.577 | -36.4199
26.6451 | -35.5773 | -36.4249
26.6512 | -35.5777 | -36.4301
26.6573 | -35.5779 | -36.435
26.6634 | -35.5782 | -36.44
26.6695 | -35.5786 | -36.4446
26.6756 | -35.5789 | -36.4498
26.6817 | -35.5792 | -36.4549
26.6878 | -35.5795 | -36.46
26.6939 | -35.5798 | -36.465
26.7 | -35.5801 | -36.4696
26.7061 | -35.5804 | -36.4745
26.7122 | -35.5807 | -36.4797
26.7183 | -35.5809 | -36.4846
26.7245 | -35.5812 | -36.4897
26.7306 | -35.5815 | -36.4946
26.7367 | -35.5818 | -36.4993
26.7428 | -35.5822 | -36.5044
26.7489 | -35.5825 | -36.5092
26.755 | -35.5828 | -36.5141
26.7611 | -35.5831 | -36.519
26.7672 | -35.5833 | -36.5242
26.7733 | -35.5836 | -36.5293
26.7794 | -35.584 | -36.5342
26.7855 | -35.5843 | -36.5392
26.7916 | -35.5846 | -36.5443
26.7977 | -35.5849 | -36.5492
26.8038 | -35.5852 | -36.5542
26.8099 | -35.5855 | -36.5592
26.816 | -35.5858 | -36.5642
26.8221 | -35.5861 | -36.5691
26.8282 | -35.5864 | -36.5741
26.8343 | -35.5867 | -36.5789
26.8404 | -35.587 | -36.5843
26.8465 | -35.5873 | -36.5893
26.8526 | -35.5876 | -36.5944
26.8587 | -35.5879 | -36.5994
26.8649 | -35.5883 | -36.6042
26.871 | -35.5886 | -36.609
26.8771 | -35.5889 | -36.6139
26.8832 | -35.5892 | -36.6189
26.8893 | -35.5896 | -36.6237
26.8954 | -35.5898 | -36.6286
26.9015 | -35.5901 | -36.6334
26.9076 | -35.5904 | -36.638
26.9137 | -35.5907 | -36.6427
26.9198 | -35.591 | -36.6481
26.9259 | -35.5912 | -36.653
26.932 | -35.5915 | -36.6579
26.9381 | -35.5918 | -36.663
26.9442 | -35.592 | -36.668
26.9503 | -35.5924 | -36.6728
26.9564 | -35.5926 | -36.6775
26.9625 | -35.5929 | -36.6825
26.9686 | -35.5933 | -36.6874
26.9747 | -35.5936 | -36.6921
26.9808 | -35.5939 | -36.6973
26.9869 | -35.5942 | -36.7022
26.993 | -35.5945 | -36.7072
26.9991 | -35.5947 | -36.712
27.0052 | -35.5951 | -36.717
27.0114 | -35.5954 | -36.7222
27.0175 | -35.5957 | -36.7272
27.0236 | -35.596 | -36.7322
27.0297 | -35.5963 | -36.7371
27.0358 | -35.5966 | -36.7422
27.0419 | -35.5969 | -36.7472
27.048 | -35.5972 | -36.752
27.0541 | -35.5976 | -36.757
27.0602 | -35.5978 | -36.7621
27.0663 | -35.5981 | -36.767
27.0724 | -35.5985 | -36.7719
27.0785 | -35.5987 | -36.7771
27.0846 | -35.5991 | -36.7819
27.0907 | -35.5994 | -36.7868
27.0968 | -35.5996 | -36.7919
27.1029 | -35.6 | -36.7969
27.109 | -35.6003 | -36.8016
27.1151 | -35.6006 | -36.8064
27.1212 | -35.6009 | -36.8114
27.1273 | -35.6011 | -36.8163
27.1334 | -35.6015 | -36.8212
27.1395 | -35.6018 | -36.826
27.1456 | -35.6021 | -36.8309
27.1518 | -35.6024 | -36.836
27.1579 | -35.6026 | -36.8408
27.164 | -35.603 | -36.8458
27.1701 | -35.6033 | -36.8504
27.1762 | -35.6036 | -36.8552
27.1823 | -35.6038 | -36.8601
27.1884 | -35.6042 | -36.8653
27.1945 | -35.6044 | -36.8702
27.2006 | -35.6047 | -36.875
27.2067 | -35.605 | -36.8802
27.2128 | -35.6053 | -36.8852
27.2189 | -35.6056 | -36.89
27.225 | -35.6059 | -36.895
27.2311 | -35.6062 | -36.9001
27.2372 | -35.6064 | -36.905
27.2433 | -35.6067 | -36.9098
27.2494 | -35.607 | -36.9151
27.2555 | -35.6073 | -36.92
27.2616 | -35.6076 | -36.9251
27.2677 | -35.6079 | -36.9301
27.2738 | -35.6081 | -36.9352
27.2799 | -35.6085 | -36.9403
27.286 | -35.6087 | -36.9449
27.2921 | -35.6091 | -36.9498
27.2983 | -35.6094 | -36.9545
27.3044 | -35.6097 | -36.9595
27.3105 | -35.6099 | -36.9645
27.3166 | -35.6102 | -36.9696
27.3227 | -35.6105 | -36.9744
27.3288 | -35.6108 | -36.9794
27.3349 | -35.6111 | -36.9845
27.341 | -35.6114 | -36.9895
27.3471 | -35.6117 | -36.9945
27.3532 | -35.612 | -36.9996
27.3593 | -35.6123 | -37.0045
27.3654 | -35.6126 | -37.0092
27.3715 | -35.6129 | -37.0142
27.3776 | -35.6132 | -37.0195
27.3837 | -35.6134 | -37.0244
27.3898 | -35.6137 | -37.0293
27.3959 | -35.614 | -37.034
27.402 | -35.6143 | -37.0386
27.4081 | -35.6146 | -37.0435
27.4142 | -35.6149 | -37.0485
27.4203 | -35.6152 | -37.0534
27.4264 | -35.6155 | -37.0585
27.4325 | -35.6158 | -37.0633
27.4387 | -35.616 | -37.0681
27.4448 | -35.6163 | -37.073
27.4509 | -35.6166 | -37.0783
27.457 | -35.6169 | -37.0831
27.4631 | -35.6172 | -37.0878
27.4692 | -35.6176 | -37.0928
27.4753 | -35.6179 | -37.0975
27.4814 | -35.6181 | -37.1026
27.4875 | -35.6184 | -37.1075
27.4936 | -35.6187 | -37.112
27.4997 | -35.619 | -37.117
27.5058 | -35.6193 | -37.1221
27.5119 | -35.6196 | -37.1271
27.518 | -35.6198 | -37.1317
27.5241 | -35.6201 | -37.1365
27.5302 | -35.6204 | -37.1414
27.5363 | -35.6206 | -37.1466
27.5424 | -35.6209 | -37.151
27.5485 | -35.6212 | -37.1562
27.5546 | -35.6215 | -37.1611
27.5607 | -35.6218 | -37.1659
27.5668 | -35.622 | -37.1709
27.5729 | -35.6223 | -37.1761
27.5791 | -35.6226 | -37.1811
27.5852 | -35.6229 | -37.1861
27.5913 | -35.6232 | -37.1914
27.5974 | -35.6235 | -37.1963
27.6035 | -35.6238 | -37.201
27.6096 | -35.6241 | -37.2059
27.6157 | -35.6244 | -37.2106
27.6218 | -35.6246 | -37.2157
27.6279 | -35.6249 | -37.2206
27.634 | -35.6252 | -37.2256
27.6401 | -35.6254 | -37.2306
27.6462 | -35.6257 | -37.2357
27.6523 | -35.6261 | -37.2404
27.6584 | -35.6263 | -37.2452
27.6645 | -35.6267 | -37.2506
27.6706 | -35.627 | -37.2554
27.6767 | -35.6273 | -37.2602
27.6828 | -35.6276 | -37.2651
27.6889 | -35.6279 | -37.2701
27.695 | -35.6282 | -37.2751
27.7011 | -35.6285 | -37.2802
27.7072 | -35.6288 | -37.2851
27.7133 | -35.629 | -37.2899
27.7194 | -35.6293 | -37.295
27.7256 | -35.6296 | -37.2997
27.7317 | -35.6299 | -37.305
27.7378 | -35.6302 | -37.3099
27.7439 | -35.6305 | -37.3153
27.75 | -35.6308 | -37.3203
27.7561 | -35.6311 | -37.325
27.7622 | -35.6314 | -37.3301
27.7683 | -35.6316 | -37.3352
27.7744 | -35.6319 | -37.3405
27.7805 | -35.6322 | -37.3454
27.7866 | -35.6325 | -37.3505
27.7927 | -35.6328 | -37.3552
27.7988 | -35.6331 | -37.36
27.8049 | -35.6334 | -37.3648
27.811 | -35.6336 | -37.3697
27.8171 | -35.6339 | -37.3746
27.8232 | -35.6342 | -37.3795
27.8293 | -35.6345 | -37.3846
27.8354 | -35.6347 | -37.3894
27.8415 | -35.635 | -37.3946
27.8476 | -35.6353 | -37.4
27.8537 | -35.6357 | -37.4054
27.8598 | -35.6359 | -37.4106
27.866 | -35.6362 | -37.4156
27.8721 | -35.6365 | -37.4205
27.8782 | -35.6368 | -37.4255
27.8843 | -35.6372 | -37.4306
27.8904 | -35.6375 | -37.4356
27.8965 | -35.6378 | -37.4407
27.9026 | -35.638 | -37.4456
27.9087 | -35.6382 | -37.4507
27.9148 | -35.6384 | -37.4555
27.9209 | -35.6387 | -37.4603
27.927 | -35.6391 | -37.4652
27.9331 | -35.6394 | -37.4698
27.9392 | -35.6397 | -37.4749
27.9453 | -35.6399 | -37.4799
27.9514 | -35.6402 | -37.4849
27.9575 | -35.6405 | -37.4901
27.9636 | -35.6408 | -37.4959
27.9697 | -35.6411 | -37.501
27.9758 | -35.6414 | -37.5066
27.9819 | -35.6417 | -37.5118
27.988 | -35.642 | -37.5165
27.9941 | -35.6422 | -37.5218
28.0002 | -35.6426 | -37.5271
28.0063 | -35.6428 | -37.5319
28.0125 | -35.643 | -37.5364
28.0186 | -35.6433 | -37.5414
28.0247 | -35.6436 | -37.5464
28.0308 | -35.6439 | -37.5512
28.0369 | -35.6443 | -37.5567
28.043 | -35.6445 | -37.5619
28.0491 | -35.6448 | -37.5667
28.0552 | -35.6451 | -37.572
28.0613 | -35.6455 | -37.5767
28.0674 | -35.6458 | -37.5809
28.0735 | -35.6462 | -37.5849

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.2916 | -41.6768
21.4908 | -35.5406 | -34.7456
16.4948 | -34.9148 | -27.4416
12.6602 | -34.5584 | -21.222
9.7171 | -34.3427 | -16.3525
7.4582 | -34.1924 | -12.3987
5.7244 | -34.0898 | -9.1995
4.3936 | -34.0285 | -6.6089
3.3722 | -34.0064 | -4.5722
2.5883 | -34.0062 | -3.0293
1.9866 | -34.0158 | -1.7966
1.5248 | -34.0258 | -0.80657
1.1703 | -34.0421 | -0.16073
1 | -34.0397 | 0.15589
0.89822 | -34.0306 | 0.42874
0.68941 | -34.0182 | 1.0079
0.52919 | -33.9933 | 1.4584
0.40612 | -34.0146 | 2.0593
0.31171 | -34.0043 | 2.5096
0.23928 | -34.0199 | 3.2331
0.18358 | -34.0001 | 3.6596
0.14091 | -34.0163 | 4.539
0.10819 | -34.0598 | 5.699
0.083017 | -34.0724 | 7.1789
0.063705 | -34.1103 | 9.0499
0.048923 | -34.2122 | 11.5534
0.037479 | -34.3128 | 14.7453
0.028801 | -34.4906 | 18.7775
0.022078 | -34.8009 | 24.1366
0.017023 | -35.298 | 30.5492
0.013018 | -36.0953 | 38.2251
0.010014 | -37.2162 | 46.1371

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003673 | -4.6993
0.007347 | -26.9053
0.01102 | -35.5622
0.014693 | -41.0385
0.018366 | -43.8508
0.02204 | -46.2652
0.025713 | -48.78
0.029386 | -49.7561
0.033059 | -50.8758
0.036733 | -52.8394
0.040406 | -54.1378
0.044079 | -54.6415
0.047752 | -56.675
0.051426 | -58.2075
0.055099 | -58.5494
0.058772 | -59.425
0.062445 | -60.0208
0.066119 | -59.2779
0.069792 | -59.9689
0.073465 | -60.8502
0.077139 | -61.6466
0.080812 | -62.718
0.084485 | -61.4235
0.088158 | -61.4687
0.091832 | -62.2732
0.095505 | -62.9225
0.099178 | -63.0048
0.10285 | -63.2769
0.10652 | -64.1585
0.1102 | -64.5405
0.11387 | -65.4994
0.11754 | -65.4684
0.12122 | -65.8851
0.12489 | -66.1791
0.12856 | -66.6352
0.13224 | -66.7634
0.13591 | -67.5992
0.13958 | -68.299
0.14326 | -68.2621
0.14693 | -68.8158
0.1506 | -69.168
0.15428 | -69.2533
0.15795 | -69.5755
0.16162 | -70.0484
0.1653 | -70.0353
0.16897 | -70.6892
0.17264 | -70.9499
0.17632 | -70.3717
0.17999 | -70.4307
0.18366 | -70.5657
0.18734 | -71.8812
0.19101 | -72.3474
0.19468 | -72.5084
0.19836 | -71.484
0.20203 | -71.6999
0.2057 | -72.0979
0.20938 | -72.8219
0.21305 | -73.2537
0.21672 | -72.7872
0.2204 | -72.679
0.22407 | -72.4204
0.22774 | -72.2917
0.23142 | -72.4431
0.23509 | -73.201
0.23876 | -73.8837
0.24244 | -74.1528
0.24611 | -75.104
0.24978 | -75.2926
0.25345 | -74.4804
0.25713 | -74.5204
0.2608 | -74.9001
0.26448 | -75.026
0.26815 | -74.4115
0.27182 | -74.6301
0.27549 | -75.056
0.27917 | -75.6622
0.28284 | -75.5934
0.28652 | -75.3815
0.29019 | -75.5243
0.29386 | -75.1095
0.29753 | -73.7162
0.30121 | -74.0537
0.30488 | -75.835
0.30855 | -76.1894
0.31223 | -75.9042
0.3159 | -76.6114
0.31957 | -76.2351
0.32325 | -76.8892
0.32692 | -76.8224
0.33059 | -76.4709
0.33427 | -76.8412
0.33794 | -78.2513
0.34161 | -76.946
0.34529 | -76.3228
0.34896 | -76.9622
0.35263 | -77.4396
0.35631 | -78.0564
0.35998 | -78.6982
0.36365 | -78.6334
0.36733 | -77.8836
0.371 | -78.2106
0.37467 | -78.5798
0.37835 | -77.7265
0.38202 | -77.3225
0.38569 | -77.7044
0.38937 | -78.4165
0.39304 | -78.7442
0.39671 | -79.1033
0.40039 | -80.4921
0.40406 | -80.3435
0.40773 | -79.2544
0.41141 | -79.3565
0.41508 | -79.5595
0.41875 | -79.1314
0.42242 | -79.5099
0.4261 | -80.4115
0.42977 | -80.3977
0.43345 | -79.0219
0.43712 | -78.9338
0.44079 | -79.2147
0.44446 | -79.4384
0.44814 | -80.4021
0.45181 | -80.6915
0.45548 | -80.949
0.45916 | -80.5327
0.46283 | -80.4204
0.4665 | -81.0389
0.47018 | -80.6421
0.47385 | -80.2048
0.47752 | -79.3797
0.4812 | -80.0721
0.48487 | -80.7891
0.48854 | -81.0861
0.49222 | -81.1857
0.49589 | -81.1535
0.49956 | -80.8504
0.50324 | -80.8435
0.50691 | -80.9176
0.51058 | -81.3332
0.51426 | -80.3695
0.51793 | -81.1223
0.5216 | -82.0113
0.52528 | -80.2501
0.52895 | -80.7431
0.53262 | -80.9563
0.5363 | -81.0577
0.53997 | -81.2729
0.54364 | -81.4768
0.54732 | -81.3947
0.55099 | -81.2134
0.55466 | -81.4915
0.55834 | -81.6495
0.56201 | -81.3918
0.56568 | -81.6424
0.56936 | -82.4111
0.57303 | -82.8335
0.5767 | -82.9248
0.58038 | -82.7186
0.58405 | -82.515
0.58772 | -82.6875
0.5914 | -82.5619
0.59507 | -82.3684
0.59874 | -82.5195
0.60242 | -82.7049
0.60609 | -83.0902
0.60976 | -84.0809
0.61343 | -84.6184
0.61711 | -83.9632
0.62078 | -83.3023
0.62445 | -83.3669
0.62813 | -83.389
0.6318 | -83.5291
0.63548 | -84.0808
0.63915 | -83.9822
0.64282 | -83.4891
0.64649 | -83.2343
0.65017 | -83.2551
0.65384 | -83.3927
0.65751 | -83.4323
0.66119 | -83.4964
0.66486 | -83.6022
0.66853 | -83.7365
0.67221 | -83.9032
0.67588 | -83.9837
0.67955 | -84.1406
0.68323 | -84.1701
0.6869 | -83.6878
0.69057 | -83.1194
0.69425 | -83.2967
0.69792 | -83.5988
0.70159 | -83.9649
0.70527 | -84.3379
0.70894 | -84.4671
0.71261 | -84.5953
0.71629 | -84.8926
0.71996 | -85.0083
0.72363 | -84.7647
0.72731 | -84.9088
0.73098 | -85.1629
0.73465 | -85.0225
0.73833 | -84.8837
0.742 | -85.1084
0.74567 | -85.1601
0.74935 | -84.7729
0.75302 | -84.7565
0.75669 | -85.0783
0.76036 | -85.1737
0.76404 | -84.8341
0.76771 | -84.3718
0.77138 | -84.6465
0.77506 | -85.2816
0.77873 | -85.3817
0.78241 | -85.0887
0.78608 | -84.9137
0.78975 | -84.8968
0.79343 | -84.7233
0.7971 | -84.943
0.80077 | -85.738
0.80444 | -85.9946
0.80812 | -85.4841
0.81179 | -85.1952
0.81546 | -85.3789
0.81914 | -85.6892
0.82281 | -85.8635
0.82648 | -85.7389
0.83016 | -85.7394
0.83383 | -85.7772
0.8375 | -85.8383
0.84118 | -86.2206
0.84485 | -86.2771
0.84852 | -85.9182
0.8522 | -85.5521
0.85587 | -85.1496
0.85954 | -85.0222
0.86322 | -85.1714
0.86689 | -85.5784
0.87056 | -86.2001
0.87424 | -86.206
0.87791 | -85.7146
0.88158 | -85.7474
0.88526 | -86.1505
0.88893 | -86.3039
0.8926 | -85.9957
0.89628 | -85.9775
0.89995 | -86.3389
0.90362 | -86.8086
0.9073 | -87.0435
0.91097 | -86.9006
0.91464 | -86.7777
0.91832 | -86.6694
0.92199 | -86.8442
0.92566 | -86.9781
0.92934 | -86.7805
0.93301 | -86.773
0.93668 | -87.031
0.94036 | -87.1873
0.94403 | -87.2625
0.9477 | -87.2693
0.95137 | -87.217
0.95505 | -87.4177
0.95872 | -87.48
0.9624 | -87.0396
0.96607 | -86.7484
0.96974 | -86.7405
0.97342 | -86.8691
0.97709 | -87.0338
0.98076 | -87.2552
0.98443 | -87.2509
0.98811 | -87.4594
0.99178 | -87.4887
0.99545 | -87.3256
0.99913 | -87.55
1.0028 | -88.1051
1.0065 | -88.1236
1.0101 | -87.6657
1.0138 | -87.4148
1.0175 | -87.1707
1.0212 | -86.8677
1.0248 | -86.8914
1.0285 | -87.362
1.0322 | -87.9727
1.0359 | -88.2942
1.0395 | -87.9057
1.0432 | -87.5802
1.0469 | -87.6714
1.0506 | -87.9837
1.0542 | -88.0497
1.0579 | -88.0891
1.0616 | -88.1499
1.0652 | -88.2126
1.0689 | -88.254
1.0726 | -88.2671
1.0763 | -88.3169
1.0799 | -88.3175
1.0836 | -88.1717
1.0873 | -88.0614
1.091 | -88.1103
1.0946 | -88.3314
1.0983 | -88.6686
1.102 | -88.8066
1.1057 | -88.6077
1.1093 | -88.4199
1.113 | -88.528
1.1167 | -88.8785
1.1203 | -89.1773
1.124 | -89.2016
1.1277 | -89.0837
1.1314 | -89.0335
1.135 | -89.1057
1.1387 | -89.291
1.1424 | -89.3805
1.1461 | -89.1091
1.1497 | -88.6868
1.1534 | -88.5001
1.1571 | -88.5956
1.1608 | -88.7499
1.1644 | -88.7971
1.1681 | -88.6701
1.1718 | -88.4464
1.1754 | -88.4379
1.1791 | -88.8097
1.1828 | -89.2833
1.1865 | -89.5578
1.1901 | -89.6676
1.1938 | -89.695
1.1975 | -89.7118
1.2012 | -89.687
1.2048 | -89.5049
1.2085 | -89.2956
1.2122 | -89.1885
1.2159 | -89.0658
1.2195 | -88.8537
1.2232 | -88.7337
1.2269 | -88.8415
1.2305 | -89.122
1.2342 | -89.429
1.2379 | -89.5896
1.2416 | -89.6053
1.2452 | -89.6145
1.2489 | -89.6164
1.2526 | -89.5354
1.2563 | -89.3896
1.2599 | -89.243
1.2636 | -89.1427
1.2673 | -89.1139
1.2709 | -89.1237
1.2746 | -89.1458
1.2783 | -89.2211
1.282 | -89.3588
1.2856 | -89.4309
1.2893 | -89.4411
1.293 | -89.6253
1.2967 | -90.002
1.3003 | -90.2778
1.304 | -90.3037
1.3077 | -90.1434
1.3114 | -89.8748
1.315 | -89.5791
1.3187 | -89.439
1.3224 | -89.5476
1.326 | -89.7747
1.3297 | -89.8818
1.3334 | -89.7844
1.3371 | -89.7465
1.3407 | -89.9814
1.3444 | -90.1933
1.3481 | -90.1175
1.3518 | -89.9796
1.3554 | -90.0326
1.3591 | -90.1694
1.3628 | -90.1008
1.3665 | -89.8105
1.3701 | -89.5894
1.3738 | -89.6505
1.3775 | -89.9476
1.3811 | -90.2707
1.3848 | -90.4624
1.3885 | -90.4349
1.3922 | -90.3016
1.3958 | -90.1528
1.3995 | -90.0308
1.4032 | -90.1485
1.4069 | -90.5033
1.4105 | -90.7826
1.4142 | -90.8593
1.4179 | -90.8244
1.4216 | -90.7633
1.4252 | -90.734
1.4289 | -90.7658
1.4326 | -90.7497
1.4362 | -90.6525
1.4399 | -90.6415
1.4436 | -90.7253
1.4473 | -90.8094
1.4509 | -90.8825
1.4546 | -90.9556
1.4583 | -91.1217
1.462 | -91.3195
1.4656 | -91.2874
1.4693 | -90.9848
1.473 | -90.6413
1.4767 | -90.4425
1.4803 | -90.4535
1.484 | -90.616
1.4877 | -90.6691
1.4913 | -90.3741
1.495 | -89.9788
1.4987 | -89.9305
1.5024 | -90.4011
1.506 | -91.2134
1.5097 | -91.9793
1.5134 | -92.4069
1.5171 | -92.4813
1.5207 | -92.2762
1.5244 | -91.9655
1.5281 | -91.7496
1.5318 | -91.6369
1.5354 | -91.5593
1.5391 | -91.4988
1.5428 | -91.3521
1.5464 | -91.0462
1.5501 | -90.8075
1.5538 | -90.8571
1.5575 | -91.0767
1.5611 | -91.2549
1.5648 | -91.3376
1.5685 | -91.3764
1.5722 | -91.4045
1.5758 | -91.3727
1.5795 | -91.4128
1.5832 | -91.7609
1.5869 | -92.1827
1.5905 | -92.4323
1.5942 | -92.4795
1.5979 | -92.3809
1.6015 | -92.2591
1.6052 | -92.1898
1.6089 | -92.1538
1.6126 | -92.0796
1.6162 | -91.9293
1.6199 | -91.7823
1.6236 | -91.7376
1.6273 | -91.8272
1.6309 | -91.9721
1.6346 | -92.0224
1.6383 | -91.933
1.6419 | -91.7687
1.6456 | -91.6267
1.6493 | -91.5544
1.653 | -91.5285
1.6566 | -91.5272
1.6603 | -91.5724
1.664 | -91.6942
1.6677 | -91.8608
1.6713 | -92.0284
1.675 | -92.1978
1.6787 | -92.3786
1.6824 | -92.5609
1.686 | -92.7107
1.6897 | -92.7477
1.6934 | -92.6131
1.697 | -92.3821
1.7007 | -92.1967
1.7044 | -92.145
1.7081 | -92.2745
1.7117 | -92.5359
1.7154 | -92.7862
1.7191 | -92.8958
1.7228 | -92.8815
1.7264 | -92.8146
1.7301 | -92.7314
1.7338 | -92.6498
1.7375 | -92.5517
1.7411 | -92.3946
1.7448 | -92.2136
1.7485 | -92.0988
1.7521 | -92.1152
1.7558 | -92.2394
1.7595 | -92.3552
1.7632 | -92.3367
1.7668 | -92.1867
1.7705 | -92.0436
1.7742 | -92.0182
1.7779 | -92.0674
1.7815 | -92.0621
1.7852 | -91.9838
1.7889 | -91.9589
1.7926 | -92.0917
1.7962 | -92.3223
1.7999 | -92.5207
1.8036 | -92.647
1.8072 | -92.7085
1.8109 | -92.6734
1.8146 | -92.5274
1.8183 | -92.32
1.8219 | -92.1818
1.8256 | -92.242
1.8293 | -92.5235
1.833 | -92.8984
1.8366 | -93.1653
1.8403 | -93.2116
1.844 | -93.0616
1.8477 | -92.8421
1.8513 | -92.6833
1.855 | -92.6328
1.8587 | -92.6959
1.8623 | -92.8408
1.866 | -92.9868
1.8697 | -93.0707
1.8734 | -93.0882
1.877 | -93.0874
1.8807 | -93.1193
1.8844 | -93.1895
1.8881 | -93.1774
1.8917 | -92.9212
1.8954 | -92.4895
1.8991 | -92.1622
1.9027 | -92.1069
1.9064 | -92.2725
1.9101 | -92.5112
1.9138 | -92.7721
1.9174 | -93.0865
1.9211 | -93.375
1.9248 | -93.5156
1.9285 | -93.5128
1.9321 | -93.4329
1.9358 | -93.3486
1.9395 | -93.3001
1.9432 | -93.3183
1.9468 | -93.3864
1.9505 | -93.4626
1.9542 | -93.5139
1.9578 | -93.5215
1.9615 | -93.5054
1.9652 | -93.4862
1.9689 | -93.4443
1.9725 | -93.3895
1.9762 | -93.3765
1.9799 | -93.4359
1.9836 | -93.5147
1.9872 | -93.5201
1.9909 | -93.4234
1.9946 | -93.2859
1.9983 | -93.1983
2.0019 | -93.1809
2.0056 | -93.1982
2.0093 | -93.1982
2.0129 | -93.169
2.0166 | -93.1567
2.0203 | -93.2712
2.024 | -93.4961
2.0276 | -93.676
2.0313 | -93.6897
2.035 | -93.5746
2.0387 | -93.3947
2.0423 | -93.2109
2.046 | -93.0838
2.0497 | -93.0426
2.0534 | -93.0584
2.057 | -93.1145
2.0607 | -93.2372
2.0644 | -93.4751
2.068 | -93.7892
2.0717 | -93.9808
2.0754 | -93.8685
2.0791 | -93.5912
2.0827 | -93.3785
2.0864 | -93.3155
2.0901 | -93.4048
2.0938 | -93.6476
2.0974 | -94.0255
2.1011 | -94.4595
2.1048 | -94.7932
2.1085 | -94.8645
2.1121 | -94.673
2.1158 | -94.4209
2.1195 | -94.3651
2.1231 | -94.342
2.1268 | -94.3943
2.1305 | -94.49
2.1342 | -94.5841
2.1378 | -94.5901
2.1415 | -94.4831
2.1452 | -94.3046
2.1489 | -94.1175
2.1525 | -93.9864
2.1562 | -93.9494
2.1599 | -93.9969
2.1636 | -94.0554
2.1672 | -94.044
2.1709 | -93.8984
2.1746 | -93.6627
2.1782 | -93.4696
2.1819 | -93.4043
2.1856 | -93.4853
2.1893 | -93.6462
2.1929 | -93.7922
2.1966 | -93.8906
2.2003 | -93.9773
2.204 | -94.0762
2.2076 | -94.1707
2.2113 | -94.2284
2.215 | -94.2336
2.2187 | -94.2165
2.2223 | -94.2435
2.226 | -94.3543
2.2297 | -94.5387
2.2333 | -94.7018
2.237 | -94.7374
2.2407 | -94.632
2.2444 | -94.472
2.248 | -94.3562
2.2517 | -94.3316
2.2554 | -94.3989
2.2591 | -94.5273
2.2627 | -94.6633
2.2664 | -94.745
2.2701 | -94.7339
2.2738 | -94.6578
2.2774 | -94.5633
2.2811 | -94.4942
2.2848 | -94.4784
2.2884 | -94.5245
2.2921 | -94.6414
2.2958 | -94.7979
2.2995 | -94.9447
2.3031 | -95.0414
2.3068 | -95.0774
2.3105 | -95.0653
2.3142 | -95.0167
2.3178 | -94.9634
2.3215 | -94.9202
2.3252 | -94.8616
2.3288 | -94.7541
2.3325 | -94.5669
2.3362 | -94.3462
2.3399 | -94.181
2.3435 | -94.1363
2.3472 | -94.2255
2.3509 | -94.4236
2.3546 | -94.66
2.3582 | -94.8438
2.3619 | -94.9396
2.3656 | -95.0007
2.3693 | -95.1033
2.3729 | -95.2972
2.3766 | -95.5427
2.3803 | -95.7415
2.3839 | -95.7847
2.3876 | -95.6683
2.3913 | -95.4892
2.395 | -95.3631
2.3986 | -95.3577
2.4023 | -95.4446
2.406 | -95.5637
2.4097 | -95.6266
2.4133 | -95.5807
2.417 | -95.3741
2.4207 | -95.0354
2.4244 | -94.6822
2.428 | -94.4256
2.4317 | -94.3376
2.4354 | -94.4153
2.439 | -94.5771
2.4427 | -94.7543
2.4464 | -94.9038
2.4501 | -95.0262
2.4537 | -95.0782
2.4574 | -95.0365
2.4611 | -94.9085
2.4648 | -94.7145
2.4684 | -94.5067
2.4721 | -94.3562
2.4758 | -94.3188
2.4795 | -94.3883
2.4831 | -94.5136
2.4868 | -94.6338
2.4905 | -94.7291
2.4941 | -94.8254
2.4978 | -94.93
2.5015 | -95.0209
2.5052 | -95.1005
2.5088 | -95.1702
2.5125 | -95.2377
2.5162 | -95.3198
2.5199 | -95.425
2.5235 | -95.5221
2.5272 | -95.5773
2.5309 | -95.5716
2.5346 | -95.5176
2.5382 | -95.4261
2.5419 | -95.2933
2.5456 | -95.1568
2.5492 | -95.0489
2.5529 | -94.9994
2.5566 | -94.9862
2.5603 | -94.9772
2.5639 | -94.9553
2.5676 | -94.9286
2.5713 | -94.9131
2.575 | -94.9327
2.5786 | -95.0185
2.5823 | -95.1835
2.586 | -95.3906
2.5897 | -95.5799
2.5933 | -95.67
2.597 | -95.6388
2.6007 | -95.5025
2.6043 | -95.3474
2.608 | -95.2351
2.6117 | -95.1785
2.6154 | -95.1557
2.619 | -95.1636
2.6227 | -95.2173
2.6264 | -95.3349
2.6301 | -95.5002
2.6337 | -95.6547
2.6374 | -95.7444
2.6411 | -95.7603
2.6448 | -95.7388
2.6484 | -95.7625
2.6521 | -95.7941
2.6558 | -95.8264
2.6594 | -95.8486
2.6631 | -95.8418
2.6668 | -95.798
2.6705 | -95.7312
2.6741 | -95.6711
2.6778 | -95.6515
2.6815 | -95.6883
2.6852 | -95.7602
2.6888 | -95.8359
2.6925 | -95.9101
2.6962 | -95.9678
2.6998 | -96.0212
2.7035 | -96.0739
2.7072 | -96.1065
2.7109 | -96.1243
2.7145 | -96.1327
2.7182 | -96.1252
2.7219 | -96.0979
2.7256 | -96.024
2.7292 | -95.8966
2.7329 | -95.7431
2.7366 | -95.5938
2.7403 | -95.4896
2.7439 | -95.4743
2.7476 | -95.55
2.7513 | -95.6989
2.7549 | -95.8807
2.7586 | -96.0346
2.7623 | -96.1354
2.766 | -96.1836
2.7696 | -96.1815
2.7733 | -96.1718
2.777 | -96.1804
2.7807 | -96.215
2.7843 | -96.2647
2.788 | -96.3092
2.7917 | -96.297
2.7954 | -96.2387
2.799 | -96.1559
2.8027 | -96.0715
2.8064 | -96.0236
2.81 | -96.0138
2.8137 | -96.0491
2.8174 | -96.1307
2.8211 | -96.2295
2.8247 | -96.3286
2.8284 | -96.4015
2.8321 | -96.4331
2.8358 | -96.4087
2.8394 | -96.358
2.8431 | -96.3063
2.8468 | -96.2993
2.8505 | -96.3772
2.8541 | -96.5259
2.8578 | -96.7028
2.8615 | -96.8361
2.8651 | -96.8813
2.8688 | -96.8464
2.8725 | -96.7551
2.8762 | -96.6446
2.8798 | -96.5294
2.8835 | -96.4401
2.8872 | -96.3855
2.8909 | -96.3583
2.8945 | -96.3656
2.8982 | -96.3922
2.9019 | -96.4284
2.9056 | -96.4737
2.9092 | -96.5117
2.9129 | -96.5252
2.9166 | -96.5136
2.9202 | -96.4832
2.9239 | -96.4509
2.9276 | -96.4366
2.9313 | -96.443
2.9349 | -96.465
2.9386 | -96.5071
2.9423 | -96.5517
2.946 | -96.5931
2.9496 | -96.6371
2.9533 | -96.7007
2.957 | -96.8077
2.9606 | -96.9505
2.9643 | -97.1035
2.968 | -97.2469
2.9717 | -97.3695
2.9753 | -97.4658
2.979 | -97.511
2.9827 | -97.4775
2.9864 | -97.3338
2.99 | -97.1196
2.9937 | -96.9063
2.9974 | -96.753
3.0011 | -96.6922
3.0047 | -96.722
3.0084 | -96.8186
3.0121 | -96.9358
3.0157 | -97.0517
3.0194 | -97.1409
3.0231 | -97.1818
3.0268 | -97.1924
3.0304 | -97.1793
3.0341 | -97.1427
3.0378 | -97.0958
3.0415 | -97.0393
3.0451 | -96.9616
3.0488 | -96.8672
3.0525 | -96.7599
3.0562 | -96.6647
3.0598 | -96.5998
3.0635 | -96.5718
3.0672 | -96.596
3.0708 | -96.6732
3.0745 | -96.7699
3.0782 | -96.8503
3.0819 | -96.8876
3.0855 | -96.877
3.0892 | -96.8368
3.0929 | -96.7986
3.0966 | -96.785
3.1002 | -96.7942
3.1039 | -96.8042
3.1076 | -96.7996
3.1113 | -96.7832
3.1149 | -96.7672
3.1186 | -96.7687
3.1223 | -96.7849
3.1259 | -96.8106
3.1296 | -96.8437
3.1333 | -96.8546
3.137 | -96.826
3.1406 | -96.7549
3.1443 | -96.6482
3.148 | -96.5619
3.1517 | -96.5272
3.1553 | -96.554
3.159 | -96.6307
3.1627 | -96.7402
3.1664 | -96.8813
3.17 | -97.0417
3.1737 | -97.2101
3.1774 | -97.3263
3.181 | -97.436
3.1847 | -97.5138
3.1884 | -97.5587
3.1921 | -97.571
3.1957 | -97.563
3.1994 | -97.5218
3.2031 | -97.4707
3.2068 | -97.4272
3.2104 | -97.4004
3.2141 | -97.4013
3.2178 | -97.4246
3.2215 | -97.4794
3.2251 | -97.5448
3.2288 | -97.6142
3.2325 | -97.6882
3.2361 | -97.7633
3.2398 | -97.8386
3.2435 | -97.9027
3.2472 | -97.9381
3.2508 | -97.9407
3.2545 | -97.9221
3.2582 | -97.8839
3.2619 | -97.847
3.2655 | -97.831
3.2692 | -97.8435
3.2729 | -97.8828
3.2766 | -97.9346
3.2802 | -97.9896
3.2839 | -98.0306
3.2876 | -98.0363
3.2912 | -97.9968
3.2949 | -97.9125
3.2986 | -97.8079
3.3023 | -97.7166
3.3059 | -97.668
3.3096 | -97.6774
3.3133 | -97.7498
3.317 | -97.8715
3.3206 | -98.0181
3.3243 | -98.1429
3.328 | -98.2164
3.3316 | -98.2242
3.3353 | -98.1863
3.339 | -98.1123
3.3427 | -98.0065
3.3463 | -97.9034
3.35 | -97.8133
3.3537 | -97.7436
3.3574 | -97.7113
3.361 | -97.7173
3.3647 | -97.7608
3.3684 | -97.8256
3.3721 | -97.876
3.3757 | -97.9039
3.3794 | -97.922
3.3831 | -97.9449
3.3867 | -97.992
3.3904 | -98.0493
3.3941 | -98.1022
3.3978 | -98.1395
3.4014 | -98.135
3.4051 | -98.0971
3.4088 | -98.0466
3.4125 | -97.9983
3.4161 | -97.9781
3.4198 | -97.988
3.4235 | -98.0368
3.4272 | -98.1277
3.4308 | -98.225
3.4345 | -98.3016
3.4382 | -98.3432
3.4418 | -98.3447
3.4455 | -98.3229
3.4492 | -98.2928
3.4529 | -98.2478
3.4565 | -98.1976
3.4602 | -98.1526
3.4639 | -98.1174
3.4676 | -98.0893
3.4712 | -98.0669
3.4749 | -98.043
3.4786 | -98.0087
3.4823 | -97.9627
3.4859 | -97.9289
3.4896 | -97.9
3.4933 | -97.8555
3.4969 | -97.7857
3.5006 | -97.6928
3.5043 | -97.5947
3.508 | -97.5145
3.5116 | -97.4675
3.5153 | -97.4564
3.519 | -97.4798
3.5227 | -97.5331
3.5263 | -97.617
3.53 | -97.7222
3.5337 | -97.8398
3.5374 | -97.9724
3.541 | -98.102
3.5447 | -98.2392
3.5484 | -98.3823
3.552 | -98.522
3.5557 | -98.6393
3.5594 | -98.7279
3.5631 | -98.7829
3.5667 | -98.8122
3.5704 | -98.8304
3.5741 | -98.8411
3.5778 | -98.8419
3.5814 | -98.8236
3.5851 | -98.7935
3.5888 | -98.7395
3.5925 | -98.6477
3.5961 | -98.5382
3.5998 | -98.4264
3.6035 | -98.3485
3.6071 | -98.3277
3.6108 | -98.3778
3.6145 | -98.4844
3.6182 | -98.6173
3.6218 | -98.7418
3.6255 | -98.8365
3.6292 | -98.8887
3.6329 | -98.8994
3.6365 | -98.8792
3.6402 | -98.8276
3.6439 | -98.7572
3.6475 | -98.6748
3.6512 | -98.5959
3.6549 | -98.5284
3.6586 | -98.4753
3.6622 | -98.4512
3.6659 | -98.4554
3.6696 | -98.4866
3.6733 | -98.5512
3.6769 | -98.6206
3.6806 | -98.6677
3.6843 | -98.6797
3.688 | -98.6494
3.6916 | -98.5775
3.6953 | -98.4857
3.699 | -98.3735
3.7026 | -98.2596
3.7063 | -98.1598
3.71 | -98.0661
3.7137 | -97.9914
3.7173 | -97.9476
3.721 | -97.9347
3.7247 | -97.9587
3.7284 | -98.0152
3.732 | -98.0915
3.7357 | -98.1923
3.7394 | -98.2932
3.7431 | -98.3839
3.7467 | -98.4558
3.7504 | -98.5047
3.7541 | -98.533
3.7577 | -98.552
3.7614 | -98.5521
3.7651 | -98.5461
3.7688 | -98.5374
3.7724 | -98.5398
3.7761 | -98.5435
3.7798 | -98.545
3.7835 | -98.5488
3.7871 | -98.556
3.7908 | -98.5663
3.7945 | -98.5686
3.7982 | -98.5707
3.8018 | -98.5811
3.8055 | -98.595
3.8092 | -98.6114
3.8128 | -98.6334
3.8165 | -98.662
3.8202 | -98.6925
3.8239 | -98.7251
3.8275 | -98.7564
3.8312 | -98.7865
3.8349 | -98.8051
3.8386 | -98.8151
3.8422 | -98.8077
3.8459 | -98.7892
3.8496 | -98.7701
3.8533 | -98.749
3.8569 | -98.7444
3.8606 | -98.7555
3.8643 | -98.7711
3.8679 | -98.7866
3.8716 | -98.7934
3.8753 | -98.7827
3.879 | -98.7627
3.8826 | -98.7492
3.8863 | -98.75
3.89 | -98.7736
3.8937 | -98.8071
3.8973 | -98.8545
3.901 | -98.8979
3.9047 | -98.9333
3.9084 | -98.9507
3.912 | -98.9493
3.9157 | -98.9408
3.9194 | -98.9318
3.923 | -98.9197
3.9267 | -98.9239
3.9304 | -98.9443
3.9341 | -98.9764
3.9377 | -99.0255
3.9414 | -99.0889
3.9451 | -99.1516
3.9488 | -99.1977
3.9524 | -99.225
3.9561 | -99.2417
3.9598 | -99.2263
3.9635 | -99.1697
3.9671 | -99.0835
3.9708 | -98.9864
3.9745 | -98.9065
3.9781 | -98.8659
3.9818 | -98.8666
3.9855 | -98.9246
3.9892 | -99.0356
3.9928 | -99.1853
3.9965 | -99.3579
4.0002 | -99.5269
4.0039 | -99.6718
4.0075 | -99.7781
4.0112 | -99.8374
4.0149 | -99.8402
4.0186 | -99.7988
4.0222 | -99.7294
4.0259 | -99.6449
4.0296 | -99.5585
4.0332 | -99.48
4.0369 | -99.4093
4.0406 | -99.3625
4.0443 | -99.3357
4.0479 | -99.3189
4.0516 | -99.3013
4.0553 | -99.2792
4.059 | -99.258
4.0626 | -99.2411
4.0663 | -99.2405
4.07 | -99.2578
4.0736 | -99.2914
4.0773 | -99.3314
4.081 | -99.3744
4.0847 | -99.4115
4.0883 | -99.419
4.092 | -99.3897
4.0957 | -99.345
4.0994 | -99.283
4.103 | -99.2191
4.1067 | -99.184
4.1104 | -99.181
4.1141 | -99.2099
4.1177 | -99.2734
4.1214 | -99.3641
4.1251 | -99.4707
4.1287 | -99.5779
4.1324 | -99.673
4.1361 | -99.7367
4.1398 | -99.7608
4.1434 | -99.7459
4.1471 | -99.7061
4.1508 | -99.6683
4.1545 | -99.6317
4.1581 | -99.6013
4.1618 | -99.5977
4.1655 | -99.6022
4.1692 | -99.6112
4.1728 | -99.6264
4.1765 | -99.6556
4.1802 | -99.6961
4.1838 | -99.7598
4.1875 | -99.8442
4.1912 | -99.9329
4.1949 | -100.0148
4.1985 | -100.0742
4.2022 | -100.114
4.2059 | -100.1301
4.2096 | -100.1086
4.2132 | -100.0573
4.2169 | -99.9911
4.2206 | -99.9002
4.2243 | -99.7864
4.2279 | -99.6832
4.2316 | -99.5946
4.2353 | -99.5376
4.2389 | -99.5484
4.2426 | -99.5417
4.2463 | -99.5616
4.25 | -99.5986
4.2536 | -99.6575
4.2573 | -99.7393
4.261 | -99.8329
4.2647 | -99.9296
4.2683 | -100.0121
4.272 | -100.077
4.2757 | -100.1354
4.2794 | -100.161
4.283 | -100.1486
4.2867 | -100.1099
4.2904 | -100.0535
4.294 | -99.9879
4.2977 | -99.9338
4.3014 | -99.8973
4.3051 | -99.8773
4.3087 | -99.8923
4.3124 | -99.9347
4.3161 | -99.9784
4.3198 | -100.0214
4.3234 | -100.0521
4.3271 | -100.0614
4.3308 | -100.0537
4.3345 | -100.0371
4.3381 | -100.019
4.3418 | -99.9954
4.3455 | -99.97
4.3491 | -99.9624
4.3528 | -99.9679
4.3565 | -99.9772
4.3602 | -99.9896
4.3638 | -100.0061
4.3675 | -100.0291
4.3712 | -100.065
4.3749 | -100.118
4.3785 | -100.167
4.3822 | -100.2082
4.3859 | -100.2435
4.3895 | -100.2668
4.3932 | -100.2767
4.3969 | -100.283
4.4006 | -100.2855
4.4042 | -100.2834
4.4079 | -100.2839
4.4116 | -100.2891
4.4153 | -100.2995
4.4189 | -100.299
4.4226 | -100.2879
4.4263 | -100.2735
4.43 | -100.2428
4.4336 | -100.1971
4.4373 | -100.1451
4.441 | -100.0825
4.4446 | -100.0222
4.4483 | -99.9594
4.452 | -99.9066
4.4557 | -99.8732
4.4593 | -99.8517
4.463 | -99.8471
4.4667 | -99.8668
4.4704 | -99.8988
4.474 | -99.9511
4.4777 | -100.0094
4.4814 | -100.0684
4.4851 | -100.1158
4.4887 | -100.1495
4.4924 | -100.1627
4.4961 | -100.1507
4.4997 | -100.1241
4.5034 | -100.0818
4.5071 | -100.0301
4.5108 | -99.9882
4.5144 | -99.9572
4.5181 | -99.9466
4.5218 | -99.9536
4.5255 | -99.9758
4.5291 | -100.0225
4.5328 | -100.0847
4.5365 | -100.1389
4.5402 | -100.1915
4.5438 | -100.2227
4.5475 | -100.2374
4.5512 | -100.2384
4.5548 | -100.2252
4.5585 | -100.2164
4.5622 | -100.2097
4.5659 | -100.2104
4.5695 | -100.2228
4.5732 | -100.2459
4.5769 | -100.282
4.5806 | -100.323
4.5842 | -100.3604
4.5879 | -100.4008
4.5916 | -100.4408
4.5953 | -100.475
4.5989 | -100.5011
4.6026 | -100.512
4.6063 | -100.5138
4.6099 | -100.5033
4.6136 | -100.4876
4.6173 | -100.4575
4.621 | -100.417
4.6246 | -100.378
4.6283 | -100.3461
4.632 | -100.3213
4.6357 | -100.3085
4.6393 | -100.3141
4.643 | -100.3396
4.6467 | -100.3779
4.6504 | -100.4309
4.654 | -100.5019
4.6577 | -100.5742
4.6614 | -100.6332
4.665 | -100.6734
4.6687 | -100.6967
4.6724 | -100.7058
4.6761 | -100.6894
4.6797 | -100.6548
4.6834 | -100.611
4.6871 | -100.5666
4.6908 | -100.5312
4.6944 | -100.5032
4.6981 | -100.4764
4.7018 | -100.466
4.7054 | -100.4641
4.7091 | -100.4547
4.7128 | -100.4393
4.7165 | -100.4273
4.7201 | -100.4051
4.7238 | -100.372
4.7275 | -100.333
4.7312 | -100.2957
4.7348 | -100.272
4.7385 | -100.2608
4.7422 | -100.272
4.7459 | -100.3087
4.7495 | -100.361
4.7532 | -100.4195
4.7569 | -100.4793
4.7605 | -100.5188
4.7642 | -100.5364
4.7679 | -100.5121
4.7716 | -100.5068
4.7752 | -100.4938
4.7789 | -100.4787
4.7826 | -100.4715
4.7863 | -100.4656
4.7899 | -100.4539
4.7936 | -100.4554
4.7973 | -100.4716
4.801 | -100.4968
4.8046 | -100.5307
4.8083 | -100.57
4.812 | -100.6112
4.8156 | -100.6563
4.8193 | -100.7032
4.823 | -100.7394
4.8267 | -100.7758
4.8303 | -100.8141
4.834 | -100.857
4.8377 | -100.8936
4.8414 | -100.9196
4.845 | -100.9363
4.8487 | -100.9374
4.8524 | -100.9251
4.8561 | -100.9003
4.8597 | -100.8686
4.8634 | -100.8384
4.8671 | -100.8101
4.8707 | -100.7813
4.8744 | -100.765
4.8781 | -100.7747
4.8818 | -100.8012
4.8854 | -100.8429
4.8891 | -100.8972
4.8928 | -100.9492
4.8965 | -101.0051
4.9001 | -101.0551
4.9038 | -101.0797
4.9075 | -101.0879
4.9112 | -101.0781
4.9148 | -101.0691
4.9185 | -101.0583
4.9222 | -101.0471
4.9258 | -101.0446
4.9295 | -101.0392
4.9332 | -101.043
4.9369 | -101.059
4.9405 | -101.0752
4.9442 | -101.0909
4.9479 | -101.0995
4.9516 | -101.0907
4.9552 | -101.0583
4.9589 | -101.0155
4.9626 | -100.966
4.9663 | -100.9173
4.9699 | -100.8753
4.9736 | -100.8392
4.9773 | -100.8271
4.9809 | -100.8376
4.9846 | -100.863
4.9883 | -100.8927
4.992 | -100.9189
4.9956 | -100.9395
4.9993 | -100.9471
5.003 | -100.94
5.0067 | -100.9206
5.0103 | -100.8812
5.014 | -100.8374
5.0177 | -100.8006
5.0214 | -100.7732
5.025 | -100.7631
5.0287 | -100.7693
5.0324 | -100.7979
5.036 | -100.8492
5.0397 | -100.9181
5.0434 | -100.9985
5.0471 | -101.0938
5.0507 | -101.1914
5.0544 | -101.2839
5.0581 | -101.3698
5.0618 | -101.4362
5.0654 | -101.4961
5.0691 | -101.5445
5.0728 | -101.5799
5.0765 | -101.6055
5.0801 | -101.6105
5.0838 | -101.6034
5.0875 | -101.588
5.0911 | -101.5486
5.0948 | -101.5011
5.0985 | -101.447
5.1022 | -101.3876
5.1058 | -101.3371
5.1095 | -101.2938
5.1132 | -101.2515
5.1169 | -101.2201
5.1205 | -101.1964
5.1242 | -101.1737
5.1279 | -101.1624
5.1315 | -101.1541
5.1352 | -101.1531
5.1389 | -101.1699
5.1426 | -101.1893
5.1462 | -101.2046
5.1499 | -101.2124
5.1536 | -101.2042
5.1573 | -101.1819
5.1609 | -101.1508
5.1646 | -101.1112
5.1683 | -101.077
5.172 | -101.0487
5.1756 | -101.0388
5.1793 | -101.0502
5.183 | -101.07
5.1866 | -101.0974
5.1903 | -101.1396
5.194 | -101.1747
5.1977 | -101.2012
5.2013 | -101.2148
5.205 | -101.2116
5.2087 | -101.197
5.2124 | -101.1646
5.216 | -101.118
5.2197 | -101.0708
5.2234 | -101.0161
5.2271 | -100.976
5.2307 | -100.9504
5.2344 | -100.9349
5.2381 | -100.9268
5.2417 | -100.9288
5.2454 | -100.943
5.2491 | -100.9626
5.2528 | -100.9831
5.2564 | -101.0012
5.2601 | -101.0233
5.2638 | -101.0426
5.2675 | -101.0509
5.2711 | -101.0533
5.2748 | -101.0553
5.2785 | -101.049
5.2822 | -101.0386
5.2858 | -101.0283
5.2895 | -101.0192
5.2932 | -101.0122
5.2968 | -101.0041
5.3005 | -100.9986
5.3042 | -101.0055
5.3079 | -101.0037
5.3115 | -101.0053
5.3152 | -101.0132
5.3189 | -101.0281
5.3226 | -101.0545
5.3262 | -101.0795
5.3299 | -101.1094
5.3336 | -101.1469
5.3373 | -101.1883
5.3409 | -101.2253
5.3446 | -101.2658
5.3483 | -101.3107
5.3519 | -101.3571
5.3556 | -101.3987
5.3593 | -101.4423
5.363 | -101.4892
5.3666 | -101.5359
5.3703 | -101.5825
5.374 | -101.6263
5.3777 | -101.6697
5.3813 | -101.7046
5.385 | -101.7299
5.3887 | -101.7485
5.3924 | -101.7626
5.396 | -101.7677
5.3997 | -101.7598
5.4034 | -101.7402
5.407 | -101.7133
5.4107 | -101.6903
5.4144 | -101.6647
5.4181 | -101.6331
5.4217 | -101.6085
5.4254 | -101.5936
5.4291 | -101.5813
5.4328 | -101.5776
5.4364 | -101.5863
5.4401 | -101.6004
5.4438 | -101.6126
5.4474 | -101.6227
5.4511 | -101.6255
5.4548 | -101.6206
5.4585 | -101.6146
5.4621 | -101.6
5.4658 | -101.5775
5.4695 | -101.5499
5.4732 | -101.5219
5.4768 | -101.4996
5.4805 | -101.4788
5.4842 | -101.461
5.4879 | -101.4533
5.4915 | -101.4558
5.4952 | -101.472
5.4989 | -101.5046
5.5025 | -101.5501
5.5062 | -101.6058
5.5099 | -101.6686
5.5136 | -101.7369
5.5172 | -101.7919
5.5209 | -101.8409
5.5246 | -101.8742
5.5283 | -101.8932
5.5319 | -101.8939
5.5356 | -101.8825
5.5393 | -101.8575
5.543 | -101.8142
5.5466 | -101.7714
5.5503 | -101.7293
5.554 | -101.6838
5.5576 | -101.6448
5.5613 | -101.6156
5.565 | -101.5942
5.5687 | -101.5697
5.5723 | -101.5451
5.576 | -101.5193
5.5797 | -101.4883
5.5834 | -101.4534
5.587 | -101.428
5.5907 | -101.4085
5.5944 | -101.3969
5.5981 | -101.3939
5.6017 | -101.4072
5.6054 | -101.4247
5.6091 | -101.4616
5.6127 | -101.5154
5.6164 | -101.5735
5.6201 | -101.6296
5.6238 | -101.6884
5.6274 | -101.7433
5.6311 | -101.7751
5.6348 | -101.789
5.6385 | -101.7841
5.6421 | -101.7533
5.6458 | -101.7072
5.6495 | -101.6543
5.6532 | -101.6043
5.6568 | -101.5473
5.6605 | -101.4998
5.6642 | -101.4597
5.6678 | -101.4319
5.6715 | -101.4141
5.6752 | -101.4113
5.6789 | -101.4251
5.6825 | -101.4576
5.6862 | -101.5062
5.6899 | -101.5662
5.6936 | -101.6311
5.6972 | -101.7021
5.7009 | -101.78
5.7046 | -101.8638
5.7083 | -101.9381
5.7119 | -102.0021
5.7156 | -102.0537
5.7193 | -102.0905
5.7229 | -102.1
5.7266 | -102.0974
5.7303 | -102.0725
5.734 | -102.0326
5.7376 | -101.9912
5.7413 | -101.9508
5.745 | -101.911
5.7487 | -101.8849
5.7523 | -101.8754
5.756 | -101.8752
5.7597 | -101.8881
5.7633 | -101.9148
5.767 | -101.9482
5.7707 | -101.9949
5.7744 | -102.0368
5.778 | -102.0695
5.7817 | -102.0957
5.7854 | -102.1039
5.7891 | -102.0932
5.7927 | -102.0694
5.7964 | -102.0321
5.8001 | -101.9848
5.8038 | -101.925
5.8074 | -101.8687
5.8111 | -101.8071
5.8148 | -101.7564
5.8184 | -101.7183
5.8221 | -101.6933
5.8258 | -101.7223
5.8295 | -101.7305
5.8331 | -101.7494
5.8368 | -101.7795
5.8405 | -101.8143
5.8442 | -101.8534
5.8478 | -101.8902
5.8515 | -101.9276
5.8552 | -101.9598
5.8589 | -101.9847
5.8625 | -102.0167
5.8662 | -102.043
5.8699 | -102.0584
5.8735 | -102.0774
5.8772 | -102.0924
5.8809 | -102.1082
5.8846 | -102.127
5.8882 | -102.1545
5.8919 | -102.1834
5.8956 | -102.2064
5.8993 | -102.2362
5.9029 | -102.2666
5.9066 | -102.2899
5.9103 | -102.2956
5.914 | -102.2868
5.9176 | -102.2654
5.9213 | -102.2292
5.925 | -102.1875
5.9286 | -102.1365
5.9323 | -102.0761
5.936 | -102.0241
5.9397 | -101.9771
5.9433 | -101.9335
5.947 | -101.9015
5.9507 | -101.883
5.9544 | -101.8755
5.958 | -101.8794
5.9617 | -101.8887
5.9654 | -101.9028
5.9691 | -101.9167
5.9727 | -101.9259
5.9764 | -101.9387
5.9801 | -101.9487
5.9837 | -101.9589
5.9874 | -101.9736
5.9911 | -101.9802
5.9948 | -101.9802
5.9984 | -101.9822
6.0021 | -101.9881
6.0058 | -102.002
6.0095 | -102.0195
6.0131 | -102.0395
6.0168 | -102.0632
6.0205 | -102.0961
6.0242 | -102.1407
6.0278 | -102.1871
6.0315 | -102.2357
6.0352 | -102.2838
6.0388 | -102.3213
6.0425 | -102.3486
6.0462 | -102.3679
6.0499 | -102.3689
6.0535 | -102.3588
6.0572 | -102.3331
6.0609 | -102.2964
6.0646 | -102.2516
6.0682 | -102.197
6.0719 | -102.1406
6.0756 | -102.09
6.0793 | -102.0477
6.0829 | -102.0137
6.0866 | -101.9957
6.0903 | -101.9892
6.0939 | -101.9971
6.0976 | -102.0108
6.1013 | -102.037
6.105 | -102.077
6.1086 | -102.1217
6.1123 | -102.1676
6.116 | -102.2115
6.1197 | -102.2443
6.1233 | -102.2816
6.127 | -102.3131
6.1307 | -102.3362
6.1344 | -102.3445
6.138 | -102.3463
6.1417 | -102.3564
6.1454 | -102.364
6.149 | -102.3729
6.1527 | -102.3826
6.1564 | -102.396
6.1601 | -102.413
6.1637 | -102.4343
6.1674 | -102.4525
6.1711 | -102.4733
6.1748 | -102.5048
6.1784 | -102.5366
6.1821 | -102.5644
6.1858 | -102.581
6.1894 | -102.5904
6.1931 | -102.591
6.1968 | -102.5776
6.2005 | -102.5467
6.2041 | -102.4992
6.2078 | -102.4368
6.2115 | -102.3713
6.2152 | -102.291
6.2188 | -102.2128
6.2225 | -102.1489
6.2262 | -102.0963
6.2299 | -102.0566
6.2335 | -102.0352
6.2372 | -102.0359
6.2409 | -102.0576
6.2445 | -102.0997
6.2482 | -102.1534
6.2519 | -102.2185
6.2556 | -102.2927
6.2592 | -102.3726
6.2629 | -102.4468
6.2666 | -102.5128
6.2703 | -102.5716
6.2739 | -102.6173
6.2776 | -102.6513
6.2813 | -102.6616
6.285 | -102.6505
6.2886 | -102.6243
6.2923 | -102.5925
6.296 | -102.5603
6.2996 | -102.5327
6.3033 | -102.4995
6.307 | -102.4723
6.3107 | -102.4477
6.3143 | -102.4297
6.318 | -102.4213
6.3217 | -102.4185
6.3254 | -102.4167
6.329 | -102.4275
6.3327 | -102.4532
6.3364 | -102.4888
6.3401 | -102.5374
6.3437 | -102.5938
6.3474 | -102.6591
6.3511 | -102.737
6.3547 | -102.8181
6.3584 | -102.8914
6.3621 | -102.9622
6.3658 | -103.0265
6.3694 | -103.0872
6.3731 | -103.1303
6.3768 | -103.1589
6.3805 | -103.169
6.3841 | -103.157
6.3878 | -103.1353
6.3915 | -103.1093
6.3952 | -103.0837
6.3988 | -103.0686
6.4025 | -103.053
6.4062 | -103.0444
6.4098 | -103.0346
6.4135 | -103.0293
6.4172 | -103.0277
6.4209 | -103.0303
6.4245 | -103.0271
6.4282 | -103.0146
6.4319 | -103.0021
6.4356 | -102.986
6.4392 | -102.9715
6.4429 | -102.9537
6.4466 | -102.9324
6.4502 | -102.9184
6.4539 | -102.904
6.4576 | -102.8904
6.4613 | -102.8793
6.4649 | -102.8778
6.4686 | -102.8814
6.4723 | -102.8937
6.476 | -102.9109
6.4796 | -102.9411
6.4833 | -102.9787
6.487 | -103.016
6.4907 | -103.0523
6.4943 | -103.0925
6.498 | -103.1346
6.5017 | -103.1815
6.5053 | -103.2318
6.509 | -103.2724
6.5127 | -103.3042
6.5164 | -103.3297
6.52 | -103.34
6.5237 | -103.3465
6.5274 | -103.3445
6.5311 | -103.3296
6.5347 | -103.2994
6.5384 | -103.2568
6.5421 | -103.203
6.5458 | -103.1412
6.5494 | -103.0677
6.5531 | -102.9927
6.5568 | -102.9213
6.5604 | -102.8468
6.5641 | -102.7757
6.5678 | -102.7107
6.5715 | -102.6544
6.5751 | -102.6062
6.5788 | -102.5688
6.5825 | -102.5409
6.5862 | -102.5225
6.5898 | -102.5102
6.5935 | -102.5062
6.5972 | -102.5153
6.6009 | -102.5387
6.6045 | -102.5686
6.6082 | -102.611
6.6119 | -102.6583
6.6155 | -102.7089
6.6192 | -102.7646
6.6229 | -102.8162
6.6266 | -102.8668
6.6302 | -102.909
6.6339 | -102.9532
6.6376 | -102.9881
6.6413 | -103.0153
6.6449 | -103.0348
6.6486 | -103.0465
6.6523 | -103.0514
6.656 | -103.047
6.6596 | -103.0403
6.6633 | -103.0333
6.667 | -103.0248
6.6706 | -103.0159
6.6743 | -103.0055
6.678 | -102.9909
6.6817 | -102.9829
6.6853 | -102.9744
6.689 | -102.9668
6.6927 | -102.9593
6.6964 | -102.9537
6.7 | -102.9513
6.7037 | -102.9589
6.7074 | -102.9694
6.7111 | -102.9813
6.7147 | -102.9975
6.7184 | -103.0145
6.7221 | -103.032
6.7257 | -103.0479
6.7294 | -103.0577
6.7331 | -103.0513
6.7368 | -103.0413
6.7404 | -103.0325
6.7441 | -103.0134
6.7478 | -102.9929
6.7515 | -102.9655
6.7551 | -102.9364
6.7588 | -102.9078
6.7625 | -102.8769
6.7662 | -102.8468
6.7698 | -102.8119
6.7735 | -102.7814
6.7772 | -102.757
6.7808 | -102.7383
6.7845 | -102.7259
6.7882 | -102.7231
6.7919 | -102.725
6.7955 | -102.7321
6.7992 | -102.7481
6.8029 | -102.7725
6.8066 | -102.8064
6.8102 | -102.8437
6.8139 | -102.8831
6.8176 | -102.9233
6.8213 | -102.9631
6.8249 | -103.0006
6.8286 | -103.0391
6.8323 | -103.0706
6.8359 | -103.098
6.8396 | -103.1179
6.8433 | -103.1405
6.847 | -103.1563
6.8506 | -103.1686
6.8543 | -103.1738
6.858 | -103.1794
6.8617 | -103.1852
6.8653 | -103.1898
6.869 | -103.201
6.8727 | -103.2029
6.8763 | -103.1994
6.88 | -103.1988
6.8837 | -103.2009
6.8874 | -103.2012
6.891 | -103.2009
6.8947 | -103.2013
6.8984 | -103.1998
6.9021 | -103.1995
6.9057 | -103.2032
6.9094 | -103.2128
6.9131 | -103.22
6.9168 | -103.2229
6.9204 | -103.2302
6.9241 | -103.2295
6.9278 | -103.2287
6.9314 | -103.2272
6.9351 | -103.2292
6.9388 | -103.2268
6.9425 | -103.2226
6.9461 | -103.2171
6.9498 | -103.2185
6.9535 | -103.2224
6.9572 | -103.2276
6.9608 | -103.2296
6.9645 | -103.2386
6.9682 | -103.2557
6.9719 | -103.268
6.9755 | -103.2916
6.9792 | -103.3227
6.9829 | -103.3541
6.9865 | -103.3847
6.9902 | -103.4251
6.9939 | -103.4658
6.9976 | -103.505
7.0012 | -103.5376
7.0049 | -103.5656
7.0086 | -103.5809
7.0123 | -103.5898
7.0159 | -103.5917
7.0196 | -103.592
7.0233 | -103.5837
7.027 | -103.5647
7.0306 | -103.5491
7.0343 | -103.5333
7.038 | -103.5132
7.0416 | -103.4893
7.0453 | -103.4646
7.049 | -103.4434
7.0527 | -103.4249
7.0563 | -103.4054
7.06 | -103.3902
7.0637 | -103.3806
7.0674 | -103.3749
7.071 | -103.3637
7.0747 | -103.3541
7.0784 | -103.3525
7.0821 | -103.3468
7.0857 | -103.3366
7.0894 | -103.3325
7.0931 | -103.3332
7.0967 | -103.3247
7.1004 | -103.3206
7.1041 | -103.3154
7.1078 | -103.3122
7.1114 | -103.3087
7.1151 | -103.3024
7.1188 | -103.2912
7.1225 | -103.2828
7.1261 | -103.2785
7.1298 | -103.2793
7.1335 | -103.2845
7.1372 | -103.2859
7.1408 | -103.2889
7.1445 | -103.2958
7.1482 | -103.3067
7.1518 | -103.3215
7.1555 | -103.3342
7.1592 | -103.3507
7.1629 | -103.3696
7.1665 | -103.39
7.1702 | -103.4122
7.1739 | -103.4338
7.1776 | -103.4566
7.1812 | -103.4792
7.1849 | -103.5038
7.1886 | -103.5286
7.1922 | -103.5572
7.1959 | -103.5807
7.1996 | -103.6015
7.2033 | -103.6229
7.2069 | -103.6408
7.2106 | -103.6576
7.2143 | -103.6751
7.218 | -103.6866
7.2216 | -103.6936
7.2253 | -103.6925
7.229 | -103.6842
7.2327 | -103.6757
7.2363 | -103.6672
7.24 | -103.6593
7.2437 | -103.6553
7.2473 | -103.6524
7.251 | -103.6547
7.2547 | -103.6596
7.2584 | -103.6641
7.262 | -103.6743
7.2657 | -103.6874
7.2694 | -103.7049
7.2731 | -103.7251
7.2767 | -103.7487
7.2804 | -103.7701
7.2841 | -103.802
7.2878 | -103.832
7.2914 | -103.8537
7.2951 | -103.8794
7.2988 | -103.8979
7.3024 | -103.9081
7.3061 | -103.915
7.3098 | -103.9123
7.3135 | -103.8952
7.3171 | -103.8684
7.3208 | -103.8309
7.3245 | -103.7901
7.3282 | -103.7495
7.3318 | -103.7104
7.3355 | -103.6674
7.3392 | -103.6189
7.3429 | -103.5741
7.3465 | -103.5318
7.3502 | -103.4964
7.3539 | -103.4658
7.3575 | -103.4392
7.3612 | -103.4187
7.3649 | -103.4022
7.3686 | -103.3906
7.3722 | -103.3856
7.3759 | -103.3856
7.3796 | -103.3914
7.3833 | -103.4074
7.3869 | -103.4341
7.3906 | -103.4666
7.3943 | -103.5041
7.398 | -103.5427
7.4016 | -103.5788
7.4053 | -103.6188
7.409 | -103.66
7.4126 | -103.6922
7.4163 | -103.6945
7.42 | -103.71
7.4237 | -103.725
7.4273 | -103.7279
7.431 | -103.7243
7.4347 | -103.7199
7.4384 | -103.7107
7.442 | -103.7024
7.4457 | -103.6967
7.4494 | -103.6922
7.4531 | -103.6894
7.4567 | -103.6932
7.4604 | -103.6927
7.4641 | -103.6867
7.4677 | -103.6883
7.4714 | -103.6915
7.4751 | -103.6926
7.4788 | -103.6937
7.4824 | -103.6896
7.4861 | -103.6866
7.4898 | -103.6876
7.4935 | -103.6887
7.4971 | -103.6928
7.5008 | -103.6921
7.5045 | -103.6989
7.5081 | -103.7131
7.5118 | -103.7287
7.5155 | -103.7387
7.5192 | -103.7555
7.5228 | -103.772
7.5265 | -103.7941
7.5302 | -103.817
7.5339 | -103.8314
7.5375 | -103.8428
7.5412 | -103.8546
7.5449 | -103.8671
7.5486 | -103.8729
7.5522 | -103.8688
7.5559 | -103.8549
7.5596 | -103.8405
7.5632 | -103.8204
7.5669 | -103.7938
7.5706 | -103.7557
7.5743 | -103.7185
7.5779 | -103.6891
7.5816 | -103.6544
7.5853 | -103.6163
7.589 | -103.5834
7.5926 | -103.5483
7.5963 | -103.5222
7.6 | -103.5034
7.6037 | -103.4842
7.6073 | -103.4664
7.611 | -103.4569
7.6147 | -103.4517
7.6183 | -103.4447
7.622 | -103.4407
7.6257 | -103.4402
7.6294 | -103.4443
7.633 | -103.4441
7.6367 | -103.447
7.6404 | -103.4498
7.6441 | -103.453
7.6477 | -103.4613
7.6514 | -103.4717
7.6551 | -103.4815
7.6588 | -103.4882
7.6624 | -103.4954
7.6661 | -103.5094
7.6698 | -103.5283
7.6734 | -103.5472
7.6771 | -103.5705
7.6808 | -103.5967
7.6845 | -103.629
7.6881 | -103.6612
7.6918 | -103.6877
7.6955 | -103.7125
7.6992 | -103.7405
7.7028 | -103.7671
7.7065 | -103.7923
7.7102 | -103.8102
7.7139 | -103.8181
7.7175 | -103.8296
7.7212 | -103.8378
7.7249 | -103.8362
7.7285 | -103.8372
7.7322 | -103.8441
7.7359 | -103.8506
7.7396 | -103.854
7.7432 | -103.8571
7.7469 | -103.8675
7.7506 | -103.8783
7.7543 | -103.8901
7.7579 | -103.9075
7.7616 | -103.9247
7.7653 | -103.9416
7.769 | -103.9613
7.7726 | -103.9833
7.7763 | -104.0031
7.78 | -104.0184
7.7836 | -104.0273
7.7873 | -104.0346
7.791 | -104.0414
7.7947 | -104.0439
7.7983 | -104.041
7.802 | -104.0317
7.8057 | -104.0212
7.8094 | -104.0029
7.813 | -103.9851
7.8167 | -103.958
7.8204 | -103.9266
7.8241 | -103.8954
7.8277 | -103.8647
7.8314 | -103.8328
7.8351 | -103.805
7.8387 | -103.7842
7.8424 | -103.7681
7.8461 | -103.7491
7.8498 | -103.7341
7.8534 | -103.7313
7.8571 | -103.7387
7.8608 | -103.746
7.8645 | -103.7542
7.8681 | -103.7639
7.8718 | -103.7738
7.8755 | -103.7816
7.8792 | -103.7837
7.8828 | -103.7832
7.8865 | -103.7764
7.8902 | -103.7617
7.8938 | -103.7524
7.8975 | -103.7372
7.9012 | -103.7123
7.9049 | -103.6835
7.9085 | -103.6569
7.9122 | -103.636
7.9159 | -103.6225
7.9196 | -103.6118
7.9232 | -103.6123
7.9269 | -103.6154
7.9306 | -103.6293
7.9342 | -103.6531
7.9379 | -103.6845
7.9416 | -103.7309
7.9453 | -103.7839
7.9489 | -103.8518
7.9526 | -103.9112
7.9563 | -103.9724
7.96 | -104.033
7.9636 | -104.098
7.9673 | -104.161
7.971 | -104.2207
7.9747 | -104.2752
7.9783 | -104.3281
7.982 | -104.3811
7.9857 | -104.4304
7.9893 | -104.472
7.993 | -104.5072
7.9967 | -104.5357
8.0004 | -104.5603
8.004 | -104.5827
8.0077 | -104.6022
8.0114 | -104.6045
8.0151 | -104.6035
8.0187 | -104.6056
8.0224 | -104.6028
8.0261 | -104.5911
8.0298 | -104.5717
8.0334 | -104.5461
8.0371 | -104.5166
8.0408 | -104.4847
8.0444 | -104.4592
8.0481 | -104.4365
8.0518 | -104.4095
8.0555 | -104.3777
8.0591 | -104.3425
8.0628 | -104.3109
8.0665 | -104.2824
8.0702 | -104.2529
8.0738 | -104.2267
8.0775 | -104.2038
8.0812 | -104.1801
8.0849 | -104.1576
8.0885 | -104.1415
8.0922 | -104.1268
8.0959 | -104.115
8.0995 | -104.1061
8.1032 | -104.1012
8.1069 | -104.0984
8.1106 | -104.1024
8.1142 | -104.1154
8.1179 | -104.1306
8.1216 | -104.1427
8.1253 | -104.1556
8.1289 | -104.1714
8.1326 | -104.1921
8.1363 | -104.2128
8.14 | -104.23
8.1436 | -104.2493
8.1473 | -104.2698
8.151 | -104.2889
8.1546 | -104.3039
8.1583 | -104.3248
8.162 | -104.342
8.1657 | -104.3451
8.1693 | -104.3492
8.173 | -104.3505
8.1767 | -104.3486
8.1804 | -104.3419
8.184 | -104.3301
8.1877 | -104.3129
8.1914 | -104.29
8.1951 | -104.2572
8.1987 | -104.2211
8.2024 | -104.1841
8.2061 | -104.1469
8.2097 | -104.1086
8.2134 | -104.0702
8.2171 | -104.0356
8.2208 | -104.0107
8.2244 | -103.9935
8.2281 | -103.9861
8.2318 | -103.9886
8.2355 | -103.9953
8.2391 | -104.0131
8.2428 | -104.0373
8.2465 | -104.0654
8.2501 | -104.0962
8.2538 | -104.1302
8.2575 | -104.1669
8.2612 | -104.2037
8.2648 | -104.239
8.2685 | -104.2708
8.2722 | -104.3001
8.2759 | -104.3329
8.2795 | -104.3555
8.2832 | -104.3756
8.2869 | -104.4003
8.2906 | -104.4149
8.2942 | -104.4244
8.2979 | -104.4305
8.3016 | -104.431
8.3052 | -104.4183
8.3089 | -104.4102
8.3126 | -104.4039
8.3163 | -104.395
8.3199 | -104.3803
8.3236 | -104.3664
8.3273 | -104.3559
8.331 | -104.3412
8.3346 | -104.3286
8.3383 | -104.3192
8.342 | -104.3081
8.3457 | -104.2943
8.3493 | -104.2784
8.353 | -104.2622
8.3567 | -104.2496
8.3603 | -104.2392
8.364 | -104.2251
8.3677 | -104.2059
8.3714 | -104.1904
8.375 | -104.1762
8.3787 | -104.1645
8.3824 | -104.1591
8.3861 | -104.152
8.3897 | -104.1454
8.3934 | -104.1431
8.3971 | -104.1399
8.4008 | -104.1417
8.4044 | -104.145
8.4081 | -104.1467
8.4118 | -104.1495
8.4154 | -104.1567
8.4191 | -104.1641
8.4228 | -104.1691
8.4265 | -104.1765
8.4301 | -104.1748
8.4338 | -104.174
8.4375 | -104.1785
8.4412 | -104.1852
8.4448 | -104.1908
8.4485 | -104.1927
8.4522 | -104.2045
8.4559 | -104.2192
8.4595 | -104.2318
8.4632 | -104.2435
8.4669 | -104.2591
8.4705 | -104.2756
8.4742 | -104.2956
8.4779 | -104.3151
8.4816 | -104.3293
8.4852 | -104.338
8.4889 | -104.352
8.4926 | -104.367
8.4963 | -104.3754
8.4999 | -104.377
8.5036 | -104.3775
8.5073 | -104.3727
8.511 | -104.3705
8.5146 | -104.3687
8.5183 | -104.3602
8.522 | -104.3498
8.5256 | -104.3363
8.5293 | -104.3213
8.533 | -104.3042
8.5367 | -104.2908
8.5403 | -104.2758
8.544 | -104.2647
8.5477 | -104.2594
8.5514 | -104.259
8.555 | -104.2673
8.5587 | -104.2798
8.5624 | -104.2994
8.5661 | -104.3233
8.5697 | -104.3502
8.5734 | -104.3801
8.5771 | -104.4054
8.5807 | -104.4277
8.5844 | -104.4523
8.5881 | -104.4758
8.5918 | -104.4941
8.5954 | -104.5063
8.5991 | -104.5073
8.6028 | -104.5121
8.6065 | -104.5162
8.6101 | -104.5125
8.6138 | -104.5058
8.6175 | -104.495
8.6211 | -104.482
8.6248 | -104.4673
8.6285 | -104.4517
8.6322 | -104.4342
8.6358 | -104.4178
8.6395 | -104.4034
8.6432 | -104.3933
8.6469 | -104.3837
8.6505 | -104.3755
8.6542 | -104.3719
8.6579 | -104.3768
8.6616 | -104.3873
8.6652 | -104.4012
8.6689 | -104.4217
8.6726 | -104.446
8.6762 | -104.4778
8.6799 | -104.5143
8.6836 | -104.5474
8.6873 | -104.5821
8.6909 | -104.6199
8.6946 | -104.6612
8.6983 | -104.6989
8.702 | -104.737
8.7056 | -104.7774
8.7093 | -104.8186
8.713 | -104.8586
8.7167 | -104.897
8.7203 | -104.9338
8.724 | -104.9712
8.7277 | -105.0104
8.7313 | -105.0452
8.735 | -105.0754
8.7387 | -105.1082
8.7424 | -105.1407
8.746 | -105.1647
8.7497 | -105.1859
8.7534 | -105.2074
8.7571 | -105.2242
8.7607 | -105.2355
8.7644 | -105.2435
8.7681 | -105.2442
8.7718 | -105.2364
8.7754 | -105.2243
8.7791 | -105.2089
8.7828 | -105.1867
8.7864 | -105.1556
8.7901 | -105.1193
8.7938 | -105.0851
8.7975 | -105.0436
8.8011 | -104.9994
8.8048 | -104.9559
8.8085 | -104.9146
8.8122 | -104.877
8.8158 | -104.837
8.8195 | -104.796
8.8232 | -104.7633
8.8269 | -104.7375
8.8305 | -104.7113
8.8342 | -104.6833
8.8379 | -104.6595
8.8415 | -104.639
8.8452 | -104.6195
8.8489 | -104.6054
8.8526 | -104.5947
8.8562 | -104.5904
8.8599 | -104.5904
8.8636 | -104.5988
8.8673 | -104.6093
8.8709 | -104.6247
8.8746 | -104.6452
8.8783 | -104.6639
8.882 | -104.68
8.8856 | -104.6952
8.8893 | -104.7112
8.893 | -104.7258
8.8966 | -104.739
8.9003 | -104.7488
8.904 | -104.7576
8.9077 | -104.758
8.9113 | -104.7541
8.915 | -104.7491
8.9187 | -104.7382
8.9224 | -104.7162
8.926 | -104.6962
8.9297 | -104.6757
8.9334 | -104.6518
8.937 | -104.6251
8.9407 | -104.5984
8.9444 | -104.5764
8.9481 | -104.5543
8.9517 | -104.5306
8.9554 | -104.5115
8.9591 | -104.5
8.9628 | -104.4963
8.9664 | -104.496
8.9701 | -104.4969
8.9738 | -104.5113
8.9775 | -104.5339
8.9811 | -104.5622
8.9848 | -104.5938
8.9885 | -104.6286
8.9921 | -104.668
8.9958 | -104.7096
8.9995 | -104.7493
9.0032 | -104.7895
9.0068 | -104.8221
9.0105 | -104.8576
9.0142 | -104.8871
9.0179 | -104.915
9.0215 | -104.9359
9.0252 | -104.9481
9.0289 | -104.9536
9.0326 | -104.9548
9.0362 | -104.9471
9.0399 | -104.9345
9.0436 | -104.9205
9.0472 | -104.9025
9.0509 | -104.8865
9.0546 | -104.8694
9.0583 | -104.8468
9.0619 | -104.8227
9.0656 | -104.8017
9.0693 | -104.7838
9.073 | -104.7678
9.0766 | -104.7526
9.0803 | -104.742
9.084 | -104.7356
9.0877 | -104.7319
9.0913 | -104.729
9.095 | -104.7308
9.0987 | -104.7367
9.1023 | -104.7464
9.106 | -104.7541
9.1097 | -104.7603
9.1134 | -104.7686
9.117 | -104.7805
9.1207 | -104.7967
9.1244 | -104.8131
9.1281 | -104.8333
9.1317 | -104.8488
9.1354 | -104.8647
9.1391 | -104.8798
9.1428 | -104.8969
9.1464 | -104.9157
9.1501 | -104.9374
9.1538 | -104.9552
9.1574 | -104.9708
9.1611 | -104.9874
9.1648 | -105.0036
9.1685 | -105.0194
9.1721 | -105.03
9.1758 | -105.0399
9.1795 | -105.0454
9.1832 | -105.0462
9.1868 | -105.0488
9.1905 | -105.0531
9.1942 | -105.0579
9.1979 | -105.0602
9.2015 | -105.0609
9.2052 | -105.0619
9.2089 | -105.0609
9.2125 | -105.0635
9.2162 | -105.0652
9.2199 | -105.0653
9.2236 | -105.0611
9.2272 | -105.0574
9.2309 | -105.0508
9.2346 | -105.042
9.2383 | -105.0349
9.2419 | -105.0276
9.2456 | -105.0223
9.2493 | -105.0157
9.253 | -105.0097
9.2566 | -105.0053
9.2603 | -105.0049
9.264 | -105.0085
9.2676 | -105.0185
9.2713 | -105.0285
9.275 | -105.0356
9.2787 | -105.0432
9.2823 | -105.0593
9.286 | -105.0785
9.2897 | -105.1023
9.2934 | -105.1282
9.297 | -105.1536
9.3007 | -105.1762
9.3044 | -105.1914
9.308 | -105.2085
9.3117 | -105.2245
9.3154 | -105.2377
9.3191 | -105.2473
9.3227 | -105.2494
9.3264 | -105.2465
9.3301 | -105.2406
9.3338 | -105.2301
9.3374 | -105.215
9.3411 | -105.2004
9.3448 | -105.1829
9.3485 | -105.1578
9.3521 | -105.136
9.3558 | -105.11
9.3595 | -105.0898
9.3631 | -105.0681
9.3668 | -105.0478
9.3705 | -105.0254
9.3742 | -105.0041
9.3778 | -104.9843
9.3815 | -104.9709
9.3852 | -104.9615
9.3889 | -104.954
9.3925 | -104.9514
9.3962 | -104.9519
9.3999 | -104.9563
9.4036 | -104.9653
9.4072 | -104.9736
9.4109 | -104.9885
9.4146 | -105.0024
9.4182 | -105.0116
9.4219 | -105.0229
9.4256 | -105.0308
9.4293 | -105.0342
9.4329 | -105.0402
9.4366 | -105.0491
9.4403 | -105.0609
9.444 | -105.068
9.4476 | -105.066
9.4513 | -105.0626
9.455 | -105.0637
9.4587 | -105.0636
9.4623 | -105.0624
9.466 | -105.0581
9.4697 | -105.0525
9.4733 | -105.0482
9.477 | -105.0438
9.4807 | -105.0406
9.4844 | -105.0369
9.488 | -105.0301
9.4917 | -105.0262
9.4954 | -105.0202
9.4991 | -105.0186
9.5027 | -105.0184
9.5064 | -105.019
9.5101 | -105.0184
9.5138 | -105.019
9.5174 | -105.019
9.5211 | -105.0234
9.5248 | -105.0338
9.5284 | -105.0439
9.5321 | -105.0526
9.5358 | -105.0619
9.5395 | -105.0703
9.5431 | -105.0759
9.5468 | -105.0866
9.5505 | -105.0972
9.5542 | -105.1034
9.5578 | -105.1055
9.5615 | -105.1075
9.5652 | -105.1062
9.5689 | -105.1054
9.5725 | -105.1014
9.5762 | -105.09
9.5799 | -105.0751
9.5835 | -105.0624
9.5872 | -105.0532
9.5909 | -105.0429
9.5946 | -105.0275
9.5982 | -105.0143
9.6019 | -105.0047
9.6056 | -105.0015
9.6093 | -105.0002
9.6129 | -104.9977
9.6166 | -105.0009
9.6203 | -105.0138
9.624 | -105.0337
9.6276 | -105.0567
9.6313 | -105.0857
9.635 | -105.1188
9.6386 | -105.1538
9.6423 | -105.1895
9.646 | -105.2326
9.6497 | -105.2788
9.6533 | -105.3256
9.657 | -105.3703
9.6607 | -105.4109
9.6644 | -105.4416
9.668 | -105.4723
9.6717 | -105.4986
9.6754 | -105.5186
9.679 | -105.5298
9.6827 | -105.5303
9.6864 | -105.5269
9.6901 | -105.5139
9.6937 | -105.4938
9.6974 | -105.4695
9.7011 | -105.4418
9.7048 | -105.4015
9.7084 | -105.3651
9.7121 | -105.326
9.7158 | -105.2863
9.7195 | -105.2489
9.7231 | -105.2106
9.7268 | -105.1713
9.7305 | -105.1306
9.7341 | -105.0957
9.7378 | -105.0659
9.7415 | -105.0385
9.7452 | -105.0116
9.7488 | -104.9875
9.7525 | -104.9682
9.7562 | -104.9521
9.7599 | -104.9375
9.7635 | -104.924
9.7672 | -104.9171
9.7709 | -104.9167
9.7746 | -104.9202
9.7782 | -104.9258
9.7819 | -104.9339
9.7856 | -104.9465
9.7892 | -104.9578
9.7929 | -104.9694
9.7966 | -104.9841
9.8003 | -105.0014
9.8039 | -105.0181
9.8076 | -105.0371
9.8113 | -105.0507
9.815 | -105.067
9.8186 | -105.0791
9.8223 | -105.0853
9.826 | -105.0938
9.8297 | -105.099
9.8333 | -105.1001
9.837 | -105.0978
9.8407 | -105.0967
9.8443 | -105.0943
9.848 | -105.0946
9.8517 | -105.0896
9.8554 | -105.0879
9.859 | -105.0862
9.8627 | -105.0877
9.8664 | -105.0956
9.8701 | -105.1055
9.8737 | -105.1231
9.8774 | -105.1412
9.8811 | -105.1678
9.8848 | -105.1961
9.8884 | -105.2288
9.8921 | -105.2661
9.8958 | -105.3056
9.8994 | -105.345
9.9031 | -105.3851
9.9068 | -105.4218
9.9105 | -105.4577
9.9141 | -105.4897
9.9178 | -105.5204
9.9215 | -105.5521
9.9252 | -105.576
9.9288 | -105.5994
9.9325 | -105.6225
9.9362 | -105.6403
9.9398 | -105.6528
9.9435 | -105.6614
9.9472 | -105.6694
9.9509 | -105.6732
9.9545 | -105.6761
9.9582 | -105.6735
9.9619 | -105.6688
9.9656 | -105.6616
9.9692 | -105.6497
9.9729 | -105.6322
9.9766 | -105.6066
9.9803 | -105.5813
9.9839 | -105.5545
9.9876 | -105.5234
9.9913 | -105.4924
9.9949 | -105.4604
9.9986 | -105.4217
10.0023 | -105.3804
10.006 | -105.336
10.0096 | -105.2904
10.0133 | -105.2511
10.017 | -105.2148
10.0207 | -105.1804
10.0243 | -105.1458
10.028 | -105.118
10.0317 | -105.0942
10.0354 | -105.0703
10.039 | -105.0481
10.0427 | -105.0336
10.0464 | -105.0223
10.05 | -105.0149
10.0537 | -105.0111
10.0574 | -105.0138
10.0611 | -105.0183
10.0647 | -105.0241
10.0684 | -105.0451
10.0721 | -105.0623
10.0758 | -105.0794
10.0794 | -105.0997
10.0831 | -105.1232
10.0868 | -105.1454
10.0905 | -105.1717
10.0941 | -105.1989
10.0978 | -105.2233
10.1015 | -105.2505
10.1051 | -105.2745
10.1088 | -105.3007
10.1125 | -105.3238
10.1162 | -105.3397
10.1198 | -105.3531
10.1235 | -105.37
10.1272 | -105.3795
10.1309 | -105.3798
10.1345 | -105.3799
10.1382 | -105.3739
10.1419 | -105.367
10.1456 | -105.3612
10.1492 | -105.3518
10.1529 | -105.3403
10.1566 | -105.3224
10.1602 | -105.2975
10.1639 | -105.2664
10.1676 | -105.2378
10.1713 | -105.2094
10.1749 | -105.1799
10.1786 | -105.1491
10.1823 | -105.12
10.186 | -105.0935
10.1896 | -105.0658
10.1933 | -105.0365
10.197 | -105.0084
10.2007 | -104.9834
10.2043 | -104.96
10.208 | -104.9411
10.2117 | -104.9235
10.2153 | -104.9096
10.219 | -104.8958
10.2227 | -104.8864
10.2264 | -104.8831
10.23 | -104.8828
10.2337 | -104.8862
10.2374 | -104.8958
10.2411 | -104.9037
10.2447 | -104.9184
10.2484 | -104.9394
10.2521 | -104.9593
10.2558 | -104.9835
10.2594 | -105.0109
10.2631 | -105.0428
10.2668 | -105.0798
10.2704 | -105.1217
10.2741 | -105.1628
10.2778 | -105.2032
10.2815 | -105.2451
10.2851 | -105.2884
10.2888 | -105.3341
10.2925 | -105.3761
10.2962 | -105.4179
10.2998 | -105.4619
10.3035 | -105.503
10.3072 | -105.5404
10.3109 | -105.5725
10.3145 | -105.6042
10.3182 | -105.6262
10.3219 | -105.6409
10.3255 | -105.651
10.3292 | -105.6538
10.3329 | -105.6505
10.3366 | -105.6423
10.3402 | -105.6278
10.3439 | -105.6101
10.3476 | -105.5877
10.3513 | -105.5584
10.3549 | -105.5255
10.3586 | -105.4901
10.3623 | -105.4566
10.3659 | -105.4221
10.3696 | -105.3885
10.3733 | -105.3533
10.377 | -105.3127
10.3806 | -105.2761
10.3843 | -105.2447
10.388 | -105.2184
10.3917 | -105.1945
10.3953 | -105.1688
10.399 | -105.1462
10.4027 | -105.1307
10.4064 | -105.1174
10.41 | -105.1023
10.4137 | -105.0923
10.4174 | -105.0884
10.421 | -105.0868
10.4247 | -105.0878
10.4284 | -105.0952
10.4321 | -105.105
10.4357 | -105.116
10.4394 | -105.1293
10.4431 | -105.1497
10.4468 | -105.1715
10.4504 | -105.193
10.4541 | -105.2161
10.4578 | -105.2438
10.4615 | -105.276
10.4651 | -105.3066
10.4688 | -105.3403
10.4725 | -105.3748
10.4761 | -105.4085
10.4798 | -105.4461
10.4835 | -105.48
10.4872 | -105.5146
10.4908 | -105.5489
10.4945 | -105.5819
10.4982 | -105.614
10.5019 | -105.6392
10.5055 | -105.6642
10.5092 | -105.6861
10.5129 | -105.7049
10.5166 | -105.7218
10.5202 | -105.7361
10.5239 | -105.7505
10.5276 | -105.7657
10.5312 | -105.7708
10.5349 | -105.7684
10.5386 | -105.761
10.5423 | -105.7539
10.5459 | -105.7466
10.5496 | -105.736
10.5533 | -105.72
10.557 | -105.7026
10.5606 | -105.6873
10.5643 | -105.6694
10.568 | -105.6501
10.5717 | -105.6313
10.5753 | -105.6126
10.579 | -105.5982
10.5827 | -105.5876
10.5863 | -105.5766
10.59 | -105.5709
10.5937 | -105.5682
10.5974 | -105.581
10.601 | -105.586
10.6047 | -105.5963
10.6084 | -105.6118
10.6121 | -105.6256
10.6157 | -105.6393
10.6194 | -105.6555
10.6231 | -105.6705
10.6268 | -105.6848
10.6304 | -105.6972
10.6341 | -105.7075
10.6378 | -105.7196
10.6414 | -105.726
10.6451 | -105.734
10.6488 | -105.7413
10.6525 | -105.7491
10.6561 | -105.7569
10.6598 | -105.7636
10.6635 | -105.772
10.6672 | -105.778
10.6708 | -105.7805
10.6745 | -105.7823
10.6782 | -105.7817
10.6818 | -105.7819
10.6855 | -105.7822
10.6892 | -105.7842
10.6929 | -105.7815
10.6965 | -105.7793
10.7002 | -105.7757
10.7039 | -105.7676
10.7076 | -105.7585
10.7112 | -105.7479
10.7149 | -105.7336
10.7186 | -105.7217
10.7223 | -105.7099
10.7259 | -105.6975
10.7296 | -105.6901
10.7333 | -105.682
10.7369 | -105.6747
10.7406 | -105.6672
10.7443 | -105.6598
10.748 | -105.654
10.7516 | -105.6483
10.7553 | -105.6465
10.759 | -105.6397
10.7627 | -105.6356
10.7663 | -105.6343
10.77 | -105.6357
10.7737 | -105.6375
10.7774 | -105.6419
10.781 | -105.6472
10.7847 | -105.6515
10.7884 | -105.6618
10.792 | -105.6726
10.7957 | -105.6826
10.7994 | -105.6937
10.8031 | -105.6977
10.8067 | -105.7029
10.8104 | -105.7067
10.8141 | -105.7112
10.8178 | -105.7106
10.8214 | -105.7018
10.8251 | -105.6948
10.8288 | -105.69
10.8325 | -105.68
10.8361 | -105.6692
10.8398 | -105.6543
10.8435 | -105.6406
10.8471 | -105.6251
10.8508 | -105.6115
10.8545 | -105.5934
10.8582 | -105.5768
10.8618 | -105.5662
10.8655 | -105.5573
10.8692 | -105.5509
10.8729 | -105.5496
10.8765 | -105.5462
10.8802 | -105.5405
10.8839 | -105.5399
10.8876 | -105.5416
10.8912 | -105.5438
10.8949 | -105.5453
10.8986 | -105.55
10.9022 | -105.5497
10.9059 | -105.5484
10.9096 | -105.55
10.9133 | -105.5502
10.9169 | -105.552
10.9206 | -105.5485
10.9243 | -105.5435
10.928 | -105.5373
10.9316 | -105.5292
10.9353 | -105.5179
10.939 | -105.5075
10.9427 | -105.4924
10.9463 | -105.4735
10.95 | -105.4553
10.9537 | -105.4377
10.9573 | -105.4207
10.961 | -105.4053
10.9647 | -105.3871
10.9684 | -105.3677
10.972 | -105.3502
10.9757 | -105.3387
10.9794 | -105.3313
10.9831 | -105.3262
10.9867 | -105.3201
10.9904 | -105.3199
10.9941 | -105.3244
10.9977 | -105.3345
11.0014 | -105.3437
11.0051 | -105.3568
11.0088 | -105.3748
11.0124 | -105.3963
11.0161 | -105.4207
11.0198 | -105.4451
11.0235 | -105.4722
11.0271 | -105.499
11.0308 | -105.522
11.0345 | -105.5453
11.0382 | -105.5673
11.0418 | -105.5855
11.0455 | -105.5963
11.0492 | -105.6076
11.0528 | -105.6174
11.0565 | -105.6238
11.0602 | -105.6288
11.0639 | -105.6262
11.0675 | -105.6186
11.0712 | -105.6073
11.0749 | -105.594
11.0786 | -105.5794
11.0822 | -105.5602
11.0859 | -105.5356
11.0896 | -105.5089
11.0933 | -105.4842
11.0969 | -105.4602
11.1006 | -105.4387
11.1043 | -105.4117
11.1079 | -105.382
11.1116 | -105.3561
11.1153 | -105.3324
11.119 | -105.3144
11.1226 | -105.2984
11.1263 | -105.2977
11.13 | -105.287
11.1337 | -105.278
11.1373 | -105.2742
11.141 | -105.2783
11.1447 | -105.2835
11.1484 | -105.2939
11.152 | -105.3072
11.1557 | -105.3229
11.1594 | -105.3432
11.163 | -105.3707
11.1667 | -105.4001
11.1704 | -105.4295
11.1741 | -105.4595
11.1777 | -105.4923
11.1814 | -105.5252
11.1851 | -105.5566
11.1888 | -105.5926
11.1924 | -105.6297
11.1961 | -105.6633
11.1998 | -105.6959
11.2035 | -105.7224
11.2071 | -105.7505
11.2108 | -105.7761
11.2145 | -105.7991
11.2181 | -105.8155
11.2218 | -105.8246
11.2255 | -105.8346
11.2292 | -105.8382
11.2328 | -105.8422
11.2365 | -105.8439
11.2402 | -105.8417
11.2439 | -105.8395
11.2475 | -105.8384
11.2512 | -105.8353
11.2549 | -105.8296
11.2586 | -105.8213
11.2622 | -105.8139
11.2659 | -105.8096
11.2696 | -105.8085
11.2732 | -105.8061
11.2769 | -105.8049
11.2806 | -105.8012
11.2843 | -105.7971
11.2879 | -105.7999
11.2916 | -105.8038
11.2953 | -105.8123
11.299 | -105.8237
11.3026 | -105.8323
11.3063 | -105.84
11.31 | -105.8511
11.3137 | -105.8623
11.3173 | -105.8733
11.321 | -105.8851
11.3247 | -105.8971
11.3283 | -105.9045
11.332 | -105.9088
11.3357 | -105.9123
11.3394 | -105.9149
11.343 | -105.9111
11.3467 | -105.9066
11.3504 | -105.899
11.3541 | -105.8811
11.3577 | -105.8653
11.3614 | -105.8494
11.3651 | -105.8284
11.3688 | -105.8087
11.3724 | -105.7831
11.3761 | -105.7536
11.3798 | -105.7249
11.3834 | -105.6989
11.3871 | -105.6767
11.3908 | -105.6525
11.3945 | -105.6271
11.3981 | -105.601
11.4018 | -105.578
11.4055 | -105.5598
11.4092 | -105.5429
11.4128 | -105.5281
11.4165 | -105.518
11.4202 | -105.5103
11.4238 | -105.5052
11.4275 | -105.5034
11.4312 | -105.5063
11.4349 | -105.5117
11.4385 | -105.5163
11.4422 | -105.5207
11.4459 | -105.5295
11.4496 | -105.5414
11.4532 | -105.5457
11.4569 | -105.5515
11.4606 | -105.5601
11.4643 | -105.5643
11.4679 | -105.5667
11.4716 | -105.5711
11.4753 | -105.5736
11.4789 | -105.5767
11.4826 | -105.5791
11.4863 | -105.5789
11.49 | -105.5763
11.4936 | -105.5742
11.4973 | -105.5712
11.501 | -105.5686
11.5047 | -105.5672
11.5083 | -105.5671
11.512 | -105.5662
11.5157 | -105.5633
11.5194 | -105.5607
11.523 | -105.5534
11.5267 | -105.5511
11.5304 | -105.5537
11.534 | -105.555
11.5377 | -105.5568
11.5414 | -105.5588
11.5451 | -105.563
11.5487 | -105.568
11.5524 | -105.5726
11.5561 | -105.5789
11.5598 | -105.5872
11.5634 | -105.594
11.5671 | -105.6014
11.5708 | -105.6122
11.5745 | -105.6234
11.5781 | -105.6363
11.5818 | -105.65
11.5855 | -105.6613
11.5891 | -105.6711
11.5928 | -105.6834
11.5965 | -105.6922
11.6002 | -105.697
11.6038 | -105.7064
11.6075 | -105.7161
11.6112 | -105.7238
11.6149 | -105.7294
11.6185 | -105.7364
11.6222 | -105.7425
11.6259 | -105.7419
11.6296 | -105.7445
11.6332 | -105.7497
11.6369 | -105.7519
11.6406 | -105.7538
11.6442 | -105.7555
11.6479 | -105.7593
11.6516 | -105.7632
11.6553 | -105.7743
11.6589 | -105.7814
11.6626 | -105.791
11.6663 | -105.8013
11.67 | -105.8118
11.6736 | -105.8209
11.6773 | -105.8309
11.681 | -105.8422
11.6846 | -105.8589
11.6883 | -105.8782
11.692 | -105.8939
11.6957 | -105.9073
11.6993 | -105.9227
11.703 | -105.9405
11.7067 | -105.96
11.7104 | -105.9792
11.714 | -105.9954
11.7177 | -106.0126
11.7214 | -106.0303
11.7251 | -106.047
11.7287 | -106.0661
11.7324 | -106.0855
11.7361 | -106.1007
11.7397 | -106.1088
11.7434 | -106.1177
11.7471 | -106.1265
11.7508 | -106.1312
11.7544 | -106.1386
11.7581 | -106.1486
11.7618 | -106.1526
11.7655 | -106.1539
11.7691 | -106.1548
11.7728 | -106.1489
11.7765 | -106.1388
11.7802 | -106.1267
11.7838 | -106.1145
11.7875 | -106.0979
11.7912 | -106.0789
11.7948 | -106.0575
11.7985 | -106.0333
11.8022 | -106.0086
11.8059 | -105.9827
11.8095 | -105.9545
11.8132 | -105.9233
11.8169 | -105.8939
11.8206 | -105.8647
11.8242 | -105.8343
11.8279 | -105.8033
11.8316 | -105.772
11.8353 | -105.7431
11.8389 | -105.7158
11.8426 | -105.6912
11.8463 | -105.6719
11.8499 | -105.6543
11.8536 | -105.6403
11.8573 | -105.6262
11.861 | -105.6166
11.8646 | -105.6135
11.8683 | -105.609
11.872 | -105.6063
11.8757 | -105.6066
11.8793 | -105.6103
11.883 | -105.6169
11.8867 | -105.6234
11.8904 | -105.6296
11.894 | -105.6384
11.8977 | -105.6471
11.9014 | -105.653
11.905 | -105.6577
11.9087 | -105.6638
11.9124 | -105.6696
11.9161 | -105.6739
11.9197 | -105.6772
11.9234 | -105.6787
11.9271 | -105.6792
11.9308 | -105.6774
11.9344 | -105.6761
11.9381 | -105.6759
11.9418 | -105.6737
11.9455 | -105.6708
11.9491 | -105.6696
11.9528 | -105.6662
11.9565 | -105.6593
11.9601 | -105.6566
11.9638 | -105.6561
11.9675 | -105.6575
11.9712 | -105.6629
11.9748 | -105.6676
11.9785 | -105.6709
11.9822 | -105.6773
11.9859 | -105.6884
11.9895 | -105.7011
11.9932 | -105.7182
11.9969 | -105.737
12.0006 | -105.7591
12.0042 | -105.7859
12.0079 | -105.8096
12.0116 | -105.8349
12.0152 | -105.8635
12.0189 | -105.8936
12.0226 | -105.924
12.0263 | -105.9561
12.0299 | -105.9909
12.0336 | -106.0263
12.0373 | -106.062
12.041 | -106.0973
12.0446 | -106.1338
12.0483 | -106.1697
12.052 | -106.2049
12.0556 | -106.2393
12.0593 | -106.2755
12.063 | -106.3106
12.0667 | -106.3399
12.0703 | -106.3689
12.074 | -106.3941
12.0777 | -106.4164
12.0814 | -106.4377
12.085 | -106.455
12.0887 | -106.4655
12.0924 | -106.4693
12.0961 | -106.4709
12.0997 | -106.4635
12.1034 | -106.45
12.1071 | -106.4329
12.1107 | -106.4133
12.1144 | -106.3909
12.1181 | -106.3668
12.1218 | -106.3378
12.1254 | -106.3062
12.1291 | -106.274
12.1328 | -106.2368
12.1365 | -106.1944
12.1401 | -106.1509
12.1438 | -106.1041
12.1475 | -106.0647
12.1512 | -106.0307
12.1548 | -105.9949
12.1585 | -105.9618
12.1622 | -105.9253
12.1658 | -105.8916
12.1695 | -105.8651
12.1732 | -105.8421
12.1769 | -105.819
12.1805 | -105.7973
12.1842 | -105.7787
12.1879 | -105.766
12.1916 | -105.7469
12.1952 | -105.7266
12.1989 | -105.7127
12.2026 | -105.6969
12.2063 | -105.6807
12.2099 | -105.6671
12.2136 | -105.6556
12.2173 | -105.6463
12.2209 | -105.6384
12.2246 | -105.6314
12.2283 | -105.6269
12.232 | -105.6225
12.2356 | -105.6208
12.2393 | -105.6211
12.243 | -105.6227
12.2467 | -105.6226
12.2503 | -105.6256
12.254 | -105.6324
12.2577 | -105.6404
12.2614 | -105.6513
12.265 | -105.6638
12.2687 | -105.6783
12.2724 | -105.6954
12.276 | -105.7152
12.2797 | -105.7356
12.2834 | -105.756
12.2871 | -105.7786
12.2907 | -105.8017
12.2944 | -105.8242
12.2981 | -105.8487
12.3018 | -105.8743
12.3054 | -105.9017
12.3091 | -105.928
12.3128 | -105.9571
12.3165 | -105.9884
12.3201 | -106.015
12.3238 | -106.0348
12.3275 | -106.0567
12.3311 | -106.0811
12.3348 | -106.1003
12.3385 | -106.1211
12.3422 | -106.1367
12.3458 | -106.1465
12.3495 | -106.1562
12.3532 | -106.1677
12.3569 | -106.1756
12.3605 | -106.1782
12.3642 | -106.181
12.3679 | -106.1768
12.3716 | -106.1706
12.3752 | -106.1658
12.3789 | -106.1577
12.3826 | -106.1526
12.3862 | -106.1469
12.3899 | -106.1372
12.3936 | -106.1235
12.3973 | -106.1085
12.4009 | -106.0955
12.4046 | -106.0869
12.4083 | -106.0781
12.412 | -106.0693
12.4156 | -106.058
12.4193 | -106.0497
12.423 | -106.0457
12.4267 | -106.0443
12.4303 | -106.042
12.434 | -106.038
12.4377 | -106.0323
12.4413 | -106.0277
12.445 | -106.0263
12.4487 | -106.0263
12.4524 | -106.0261
12.456 | -106.0263
12.4597 | -106.0299
12.4634 | -106.0353
12.4671 | -106.0383
12.4707 | -106.0446
12.4744 | -106.0543
12.4781 | -106.0629
12.4817 | -106.07
12.4854 | -106.0783
12.4891 | -106.0889
12.4928 | -106.0971
12.4964 | -106.1059
12.5001 | -106.1156
12.5038 | -106.1215
12.5075 | -106.1237
12.5111 | -106.1268
12.5148 | -106.1291
12.5185 | -106.129
12.5222 | -106.1298
12.5258 | -106.1325
12.5295 | -106.133
12.5332 | -106.1303
12.5368 | -106.1319
12.5405 | -106.1295
12.5442 | -106.1228
12.5479 | -106.118
12.5515 | -106.1133
12.5552 | -106.1063
12.5589 | -106.0972
12.5626 | -106.0837
12.5662 | -106.0666
12.5699 | -106.0505
12.5736 | -106.0383
12.5773 | -106.0254
12.5809 | -106.0053
12.5846 | -105.9818
12.5883 | -105.9616
12.5919 | -105.9422
12.5956 | -105.9243
12.5993 | -105.9061
12.603 | -105.8879
12.6066 | -105.869
12.6103 | -105.854
12.614 | -105.8399
12.6177 | -105.8278
12.6213 | -105.8196
12.625 | -105.8108
12.6287 | -105.8016
12.6324 | -105.7952
12.636 | -105.7936
12.6397 | -105.7911
12.6434 | -105.7888
12.647 | -105.79
12.6507 | -105.7946
12.6544 | -105.7976
12.6581 | -105.8049
12.6617 | -105.8151
12.6654 | -105.8266
12.6691 | -105.8392
12.6728 | -105.8519
12.6764 | -105.8661
12.6801 | -105.8846
12.6838 | -105.9082
12.6875 | -105.9309
12.6911 | -105.952
12.6948 | -105.9739
12.6985 | -105.9957
12.7021 | -106.0187
12.7058 | -106.0381
12.7095 | -106.0543
12.7132 | -106.0717
12.7168 | -106.0763
12.7205 | -106.0885
12.7242 | -106.1001
12.7279 | -106.1075
12.7315 | -106.1116
12.7352 | -106.1147
12.7389 | -106.1156
12.7425 | -106.1132
12.7462 | -106.1142
12.7499 | -106.1165
12.7536 | -106.1162
12.7572 | -106.1117
12.7609 | -106.1058
12.7646 | -106.0995
12.7683 | -106.0905
12.7719 | -106.0799
12.7756 | -106.0734
12.7793 | -106.0681
12.783 | -106.06
12.7866 | -106.05
12.7903 | -106.0426
12.794 | -106.0372
12.7976 | -106.0314
12.8013 | -106.0284
12.805 | -106.0279
12.8087 | -106.0298
12.8123 | -106.0356
12.816 | -106.0415
12.8197 | -106.0452
12.8234 | -106.0506
12.827 | -106.0588
12.8307 | -106.0692
12.8344 | -106.0775
12.8381 | -106.0867
12.8417 | -106.0958
12.8454 | -106.1047
12.8491 | -106.1134
12.8527 | -106.1224
12.8564 | -106.1317
12.8601 | -106.1382
12.8638 | -106.1413
12.8674 | -106.1412
12.8711 | -106.1443
12.8748 | -106.1491
12.8785 | -106.1512
12.8821 | -106.1538
12.8858 | -106.1546
12.8895 | -106.1565
12.8932 | -106.1568
12.8968 | -106.1564
12.9005 | -106.1572
12.9042 | -106.1584
12.9078 | -106.1562
12.9115 | -106.1505
12.9152 | -106.1474
12.9189 | -106.1428
12.9225 | -106.1377
12.9262 | -106.1333
12.9299 | -106.1301
12.9336 | -106.1253
12.9372 | -106.1192
12.9409 | -106.1074
12.9446 | -106.0924
12.9483 | -106.0817
12.9519 | -106.075
12.9556 | -106.0687
12.9593 | -106.0592
12.9629 | -106.0465
12.9666 | -106.0331
12.9703 | -106.0224
12.974 | -106.0156
12.9776 | -106.0069
12.9813 | -105.9974
12.985 | -105.9929
12.9887 | -105.987
12.9923 | -105.9792
12.996 | -105.9784
12.9997 | -105.9816
13.0034 | -105.9888
13.007 | -106
13.0107 | -106.0154
13.0144 | -106.0322
13.018 | -106.0521
13.0217 | -106.0766
13.0254 | -106.1041
13.0291 | -106.1339
13.0327 | -106.1684
13.0364 | -106.2004
13.0401 | -106.2327
13.0438 | -106.2655
13.0474 | -106.2969
13.0511 | -106.3276
13.0548 | -106.3557
13.0585 | -106.3746
13.0621 | -106.3932
13.0658 | -106.4154
13.0695 | -106.435
13.0731 | -106.4488
13.0768 | -106.4591
13.0805 | -106.4699
13.0842 | -106.4758
13.0878 | -106.4728
13.0915 | -106.4632
13.0952 | -106.4549
13.0989 | -106.4482
13.1025 | -106.4315
13.1062 | -106.4117
13.1099 | -106.3887
13.1136 | -106.3615
13.1172 | -106.33
13.1209 | -106.298
13.1246 | -106.2663
13.1282 | -106.2362
13.1319 | -106.2094
13.1356 | -106.1797
13.1393 | -106.147
13.1429 | -106.1184
13.1466 | -106.092
13.1503 | -106.0694
13.154 | -106.0484
13.1576 | -106.0291
13.1613 | -106.0132
13.165 | -106.0049
13.1686 | -105.9996
13.1723 | -105.9958
13.176 | -105.9958
13.1797 | -105.9998
13.1833 | -106.0023
13.187 | -106.0077
13.1907 | -106.0148
13.1944 | -106.0271
13.198 | -106.0401
13.2017 | -106.0513
13.2054 | -106.0587
13.2091 | -106.0681
13.2127 | -106.0776
13.2164 | -106.0881
13.2201 | -106.0978
13.2237 | -106.1066
13.2274 | -106.1127
13.2311 | -106.1188
13.2348 | -106.1224
13.2384 | -106.1251
13.2421 | -106.1226
13.2458 | -106.1008
13.2495 | -106.1005
13.2531 | -106.0939
13.2568 | -106.0818
13.2605 | -106.0679
13.2642 | -106.0571
13.2678 | -106.0426
13.2715 | -106.0241
13.2752 | -106.0081
13.2788 | -105.9933
13.2825 | -105.9802
13.2862 | -105.9619
13.2899 | -105.944
13.2935 | -105.9295
13.2972 | -105.9163
13.3009 | -105.9045
13.3046 | -105.8955
13.3082 | -105.8884
13.3119 | -105.8787
13.3156 | -105.8722
13.3193 | -105.869
13.3229 | -105.8656
13.3266 | -105.8627
13.3303 | -105.86
13.3339 | -105.8548
13.3376 | -105.8516
13.3413 | -105.8508
13.345 | -105.8537
13.3486 | -105.8538
13.3523 | -105.8535
13.356 | -105.8503
13.3597 | -105.846
13.3633 | -105.8428
13.367 | -105.839
13.3707 | -105.834
13.3744 | -105.8298
13.378 | -105.8263
13.3817 | -105.8191
13.3854 | -105.8075
13.389 | -105.7981
13.3927 | -105.7896
13.3964 | -105.7816
13.4001 | -105.7701
13.4037 | -105.7605
13.4074 | -105.748
13.4111 | -105.7301
13.4148 | -105.7156
13.4184 | -105.7073
13.4221 | -105.6949
13.4258 | -105.6838
13.4294 | -105.6731
13.4331 | -105.666
13.4368 | -105.6594
13.4405 | -105.649
13.4441 | -105.6398
13.4478 | -105.6319
13.4515 | -105.6242
13.4552 | -105.6188
13.4588 | -105.6139
13.4625 | -105.6117
13.4662 | -105.6113
13.4699 | -105.6118
13.4735 | -105.6127
13.4772 | -105.6138
13.4809 | -105.6155
13.4845 | -105.6183
13.4882 | -105.6227
13.4919 | -105.6283
13.4956 | -105.6355
13.4992 | -105.6445
13.5029 | -105.6544
13.5066 | -105.669
13.5103 | -105.6837
13.5139 | -105.7001
13.5176 | -105.7148
13.5213 | -105.7289
13.525 | -105.7486
13.5286 | -105.7706
13.5323 | -105.79
13.536 | -105.8064
13.5396 | -105.8214
13.5433 | -105.8408
13.547 | -105.8591
13.5507 | -105.8749
13.5543 | -105.8888
13.558 | -105.9021
13.5617 | -105.915
13.5654 | -105.9264
13.569 | -105.9353
13.5727 | -105.9417
13.5764 | -105.9507
13.5801 | -105.9603
13.5837 | -105.9694
13.5874 | -105.9759
13.5911 | -105.9837
13.5947 | -105.9909
13.5984 | -105.9957
13.6021 | -105.9996
13.6058 | -106.008
13.6094 | -106.0168
13.6131 | -106.0192
13.6168 | -106.0249
13.6205 | -106.0276
13.6241 | -106.0322
13.6278 | -106.0389
13.6315 | -106.0463
13.6352 | -106.0539
13.6388 | -106.0578
13.6425 | -106.0627
13.6462 | -106.0697
13.6498 | -106.0787
13.6535 | -106.0845
13.6572 | -106.0902
13.6609 | -106.0959
13.6645 | -106.1
13.6682 | -106.1059
13.6719 | -106.1106
13.6756 | -106.1187
13.6792 | -106.1252
13.6829 | -106.1319
13.6866 | -106.1357
13.6903 | -106.1378
13.6939 | -106.143
13.6976 | -106.1511
13.7013 | -106.1565
13.7049 | -106.1608
13.7086 | -106.1632
13.7123 | -106.1669
13.716 | -106.1706
13.7196 | -106.1715
13.7233 | -106.1684
13.727 | -106.1673
13.7307 | -106.167
13.7343 | -106.165
13.738 | -106.162
13.7417 | -106.1591
13.7454 | -106.1596
13.749 | -106.1602
13.7527 | -106.1588
13.7564 | -106.1564
13.76 | -106.1528
13.7637 | -106.1514
13.7674 | -106.1502
13.7711 | -106.1482
13.7747 | -106.1489
13.7784 | -106.1493
13.7821 | -106.1482
13.7858 | -106.1491
13.7894 | -106.1472
13.7931 | -106.1492
13.7968 | -106.1522
13.8004 | -106.1539
13.8041 | -106.1547
13.8078 | -106.1576
13.8115 | -106.1632
13.8151 | -106.1662
13.8188 | -106.1729
13.8225 | -106.1789
13.8262 | -106.1809
13.8298 | -106.186
13.8335 | -106.1945
13.8372 | -106.204
13.8409 | -106.2132
13.8445 | -106.2213
13.8482 | -106.2295
13.8519 | -106.2381
13.8555 | -106.2479
13.8592 | -106.2608
13.8629 | -106.2726
13.8666 | -106.2823
13.8702 | -106.2929
13.8739 | -106.3006
13.8776 | -106.3094
13.8813 | -106.3167
13.8849 | -106.3224
13.8886 | -106.3269
13.8923 | -106.3277
13.896 | -106.3242
13.8996 | -106.3191
13.9033 | -106.3147
13.907 | -106.3117
13.9106 | -106.308
13.9143 | -106.3036
13.918 | -106.2948
13.9217 | -106.2893
13.9253 | -106.2798
13.929 | -106.2679
13.9327 | -106.2553
13.9364 | -106.2411
13.94 | -106.2289
13.9437 | -106.2167
13.9474 | -106.2062
13.9511 | -106.1939
13.9547 | -106.1831
13.9584 | -106.1658
13.9621 | -106.1357
13.9657 | -106.1089
13.9694 | -106.0884
13.9731 | -106.0699
13.9768 | -106.0515
13.9804 | -106.0332
13.9841 | -106.0151
13.9878 | -105.9974
13.9915 | -105.9809
13.9951 | -105.9604
13.9988 | -105.9407
14.0025 | -105.9224
14.0062 | -105.9025
14.0098 | -105.8786
14.0135 | -105.8543
14.0172 | -105.8334
14.0208 | -105.8111
14.0245 | -105.7893
14.0282 | -105.771
14.0319 | -105.7516
14.0355 | -105.7317
14.0392 | -105.7133
14.0429 | -105.6995
14.0466 | -105.6903
14.0502 | -105.6818
14.0539 | -105.6767
14.0576 | -105.6757
14.0613 | -105.6745
14.0649 | -105.6787
14.0686 | -105.6892
14.0723 | -105.7031
14.0759 | -105.7223
14.0796 | -105.7457
14.0833 | -105.774
14.087 | -105.808
14.0906 | -105.8463
14.0943 | -105.8867
14.098 | -105.9314
14.1017 | -105.9785
14.1053 | -106.0249
14.109 | -106.0726
14.1127 | -106.1199
14.1164 | -106.1698
14.12 | -106.2223
14.1237 | -106.2731
14.1274 | -106.322
14.131 | -106.3708
14.1347 | -106.4203
14.1384 | -106.4666
14.1421 | -106.5078
14.1457 | -106.5457
14.1494 | -106.582
14.1531 | -106.6171
14.1568 | -106.6482
14.1604 | -106.6843
14.1641 | -106.7251
14.1678 | -106.7526
14.1715 | -106.7683
14.1751 | -106.7788
14.1788 | -106.7864
14.1825 | -106.7934
14.1861 | -106.7974
14.1898 | -106.7989
14.1935 | -106.7986
14.1972 | -106.7986
14.2008 | -106.7967
14.2045 | -106.7939
14.2082 | -106.7902
14.2119 | -106.7864
14.2155 | -106.7829
14.2192 | -106.7772
14.2229 | -106.7677
14.2265 | -106.7581
14.2302 | -106.7489
14.2339 | -106.7397
14.2376 | -106.7295
14.2412 | -106.7195
14.2449 | -106.7056
14.2486 | -106.6931
14.2523 | -106.6811
14.2559 | -106.675
14.2596 | -106.667
14.2633 | -106.6603
14.267 | -106.653
14.2706 | -106.647
14.2743 | -106.6435
14.278 | -106.6411
14.2816 | -106.6354
14.2853 | -106.6328
14.289 | -106.6308
14.2927 | -106.6294
14.2963 | -106.629
14.3 | -106.6335
14.3037 | -106.6374
14.3074 | -106.6453
14.311 | -106.6468
14.3147 | -106.648
14.3184 | -106.6495
14.3221 | -106.6501
14.3257 | -106.6507
14.3294 | -106.6507
14.3331 | -106.6544
14.3367 | -106.6565
14.3404 | -106.6554
14.3441 | -106.6568
14.3478 | -106.6547
14.3514 | -106.6511
14.3551 | -106.6459
14.3588 | -106.639
14.3625 | -106.6311
14.3661 | -106.6177
14.3698 | -106.6015
14.3735 | -106.5866
14.3772 | -106.5692
14.3808 | -106.5474
14.3845 | -106.5256
14.3882 | -106.5017
14.3918 | -106.4774
14.3955 | -106.4493
14.3992 | -106.4174
14.4029 | -106.3866
14.4065 | -106.3566
14.4102 | -106.3258
14.4139 | -106.2958
14.4176 | -106.2669
14.4212 | -106.2375
14.4249 | -106.2119
14.4286 | -106.1934
14.4323 | -106.1738
14.4359 | -106.154
14.4396 | -106.137
14.4433 | -106.1246
14.4469 | -106.1141
14.4506 | -106.1082
14.4543 | -106.1018
14.458 | -106.1
14.4616 | -106.0971
14.4653 | -106.0966
14.469 | -106.1014
14.4727 | -106.1058
14.4763 | -106.1105
14.48 | -106.1171
14.4837 | -106.1252
14.4873 | -106.1352
14.491 | -106.1461
14.4947 | -106.1596
14.4984 | -106.1716
14.502 | -106.1855
14.5057 | -106.1984
14.5094 | -106.2117
14.5131 | -106.2248
14.5167 | -106.2374
14.5204 | -106.2524
14.5241 | -106.2662
14.5278 | -106.282
14.5314 | -106.2956
14.5351 | -106.308
14.5388 | -106.3181
14.5424 | -106.3256
14.5461 | -106.3335
14.5498 | -106.3424
14.5535 | -106.349
14.5571 | -106.3571
14.5608 | -106.3626
14.5645 | -106.3674
14.5682 | -106.3701
14.5718 | -106.3748
14.5755 | -106.3793
14.5792 | -106.3847
14.5829 | -106.3871
14.5865 | -106.388
14.5902 | -106.3893
14.5939 | -106.3903
14.5975 | -106.3891
14.6012 | -106.3872
14.6049 | -106.3905
14.6086 | -106.3961
14.6122 | -106.3949
14.6159 | -106.3902
14.6196 | -106.3867
14.6233 | -106.3833
14.6269 | -106.3794
14.6306 | -106.3777
14.6343 | -106.3746
14.638 | -106.3694
14.6416 | -106.3651
14.6453 | -106.364
14.649 | -106.3653
14.6526 | -106.3656
14.6563 | -106.3659
14.66 | -106.3686
14.6637 | -106.3716
14.6673 | -106.3724
14.671 | -106.3757
14.6747 | -106.3824
14.6784 | -106.389
14.682 | -106.3977
14.6857 | -106.4086
14.6894 | -106.4157
14.6931 | -106.4231
14.6967 | -106.4342
14.7004 | -106.4447
14.7041 | -106.4527
14.7077 | -106.463
14.7114 | -106.4732
14.7151 | -106.4814
14.7188 | -106.4887
14.7224 | -106.4959
14.7261 | -106.5045
14.7298 | -106.5117
14.7335 | -106.5162
14.7371 | -106.5175
14.7408 | -106.52
14.7445 | -106.5208
14.7482 | -106.5226
14.7518 | -106.523
14.7555 | -106.5175
14.7592 | -106.5124
14.7628 | -106.5077
14.7665 | -106.5029
14.7702 | -106.4968
14.7739 | -106.4875
14.7775 | -106.4803
14.7812 | -106.4695
14.7849 | -106.4601
14.7886 | -106.4512
14.7922 | -106.4427
14.7959 | -106.4375
14.7996 | -106.4296
14.8033 | -106.4224
14.8069 | -106.4185
14.8106 | -106.4164
14.8143 | -106.4121
14.8179 | -106.4113
14.8216 | -106.4126
14.8253 | -106.4146
14.829 | -106.4149
14.8326 | -106.4152
14.8363 | -106.413
14.84 | -106.4117
14.8437 | -106.4112
14.8473 | -106.4125
14.851 | -106.4131
14.8547 | -106.4134
14.8583 | -106.4118
14.862 | -106.4094
14.8657 | -106.4094
14.8694 | -106.4114
14.873 | -106.412
14.8767 | -106.4131
14.8804 | -106.4138
14.8841 | -106.4161
14.8877 | -106.418
14.8914 | -106.4219
14.8951 | -106.4293
14.8988 | -106.4369
14.9024 | -106.4483
14.9061 | -106.4625
14.9098 | -106.4746
14.9134 | -106.4865
14.9171 | -106.5009
14.9208 | -106.5148
14.9245 | -106.5265
14.9281 | -106.5413
14.9318 | -106.5563
14.9355 | -106.572
14.9392 | -106.5867
14.9428 | -106.5999
14.9465 | -106.6114
14.9502 | -106.6234
14.9539 | -106.6358
14.9575 | -106.65
14.9612 | -106.6622
14.9649 | -106.6709
14.9685 | -106.6839
14.9722 | -106.6924
14.9759 | -106.6995
14.9796 | -106.7043
14.9832 | -106.7052
14.9869 | -106.706
14.9906 | -106.7056
14.9943 | -106.7066
14.9979 | -106.7042
15.0016 | -106.6972
15.0053 | -106.6895
15.009 | -106.6832
15.0126 | -106.6762
15.0163 | -106.6661
15.02 | -106.6505
15.0236 | -106.6323
15.0273 | -106.6165
15.031 | -106.5996
15.0347 | -106.5835
15.0383 | -106.5637
15.042 | -106.5452
15.0457 | -106.5269
15.0494 | -106.5091
15.053 | -106.4911
15.0567 | -106.4743
15.0604 | -106.4586
15.0641 | -106.4419
15.0677 | -106.4295
15.0714 | -106.4194
15.0751 | -106.4115
15.0787 | -106.4041
15.0824 | -106.3972
15.0861 | -106.3938
15.0898 | -106.3932
15.0934 | -106.3948
15.0971 | -106.4016
15.1008 | -106.4051
15.1045 | -106.4086
15.1081 | -106.4139
15.1118 | -106.4225
15.1155 | -106.4306
15.1192 | -106.4409
15.1228 | -106.4508
15.1265 | -106.462
15.1302 | -106.4739
15.1338 | -106.4873
15.1375 | -106.5007
15.1412 | -106.5161
15.1449 | -106.5305
15.1485 | -106.5468
15.1522 | -106.5606
15.1559 | -106.5729
15.1596 | -106.5853
15.1632 | -106.598
15.1669 | -106.6098
15.1706 | -106.6178
15.1743 | -106.623
15.1779 | -106.628
15.1816 | -106.6324
15.1853 | -106.6356
15.1889 | -106.6376
15.1926 | -106.6373
15.1963 | -106.6369
15.2 | -106.6384
15.2036 | -106.6381
15.2073 | -106.6345
15.211 | -106.628
15.2147 | -106.6237
15.2183 | -106.6205
15.222 | -106.6138
15.2257 | -106.6023
15.2293 | -106.5939
15.233 | -106.5882
15.2367 | -106.5844
15.2404 | -106.5796
15.244 | -106.5731
15.2477 | -106.5638
15.2514 | -106.5571
15.2551 | -106.5531
15.2587 | -106.5466
15.2624 | -106.5391
15.2661 | -106.5295
15.2698 | -106.5175
15.2734 | -106.5053
15.2771 | -106.492
15.2808 | -106.4785
15.2844 | -106.4628
15.2881 | -106.4484
15.2918 | -106.4302
15.2955 | -106.414
15.2991 | -106.3995
15.3028 | -106.3828
15.3065 | -106.3672
15.3102 | -106.3519
15.3138 | -106.3335
15.3175 | -106.3189
15.3212 | -106.3061
15.3249 | -106.2907
15.3285 | -106.2733
15.3322 | -106.2613
15.3359 | -106.2492
15.3395 | -106.2398
15.3432 | -106.2318
15.3469 | -106.2225
15.3506 | -106.2126
15.3542 | -106.2067
15.3579 | -106.2008
15.3616 | -106.1952
15.3653 | -106.1978
15.3689 | -106.1887
15.3726 | -106.1801
15.3763 | -106.1725
15.38 | -106.167
15.3836 | -106.1637
15.3873 | -106.1587
15.391 | -106.1548
15.3946 | -106.1502
15.3983 | -106.1468
15.402 | -106.1456
15.4057 | -106.1395
15.4093 | -106.1351
15.413 | -106.1288
15.4167 | -106.1218
15.4204 | -106.115
15.424 | -106.109
15.4277 | -106.1023
15.4314 | -106.0943
15.4351 | -106.0853
15.4387 | -106.0745
15.4424 | -106.0661
15.4461 | -106.0585
15.4497 | -106.0502
15.4534 | -106.0398
15.4571 | -106.0295
15.4608 | -106.0196
15.4644 | -106.007
15.4681 | -105.9931
15.4718 | -105.9809
15.4755 | -105.9712
15.4791 | -105.9641
15.4828 | -105.9553
15.4865 | -105.9466
15.4902 | -105.9401
15.4938 | -105.9344
15.4975 | -105.9316
15.5012 | -105.9324
15.5048 | -105.9344
15.5085 | -105.9331
15.5122 | -105.9345
15.5159 | -105.9382
15.5195 | -105.942
15.5232 | -105.9464
15.5269 | -105.9512
15.5306 | -105.9605
15.5342 | -105.9724
15.5379 | -105.9807
15.5416 | -105.9885
15.5452 | -105.9994
15.5489 | -106.0094
15.5526 | -106.0177
15.5563 | -106.0275
15.5599 | -106.0351
15.5636 | -106.0421
15.5673 | -106.0532
15.571 | -106.062
15.5746 | -106.0686
15.5783 | -106.0764
15.582 | -106.0822
15.5857 | -106.0869
15.5893 | -106.0936
15.593 | -106.1024
15.5967 | -106.1119
15.6003 | -106.1155
15.604 | -106.1184
15.6077 | -106.12
15.6114 | -106.1231
15.615 | -106.1223
15.6187 | -106.1227
15.6224 | -106.1213
15.6261 | -106.1206
15.6297 | -106.1228
15.6334 | -106.1238
15.6371 | -106.123
15.6408 | -106.1199
15.6444 | -106.1147
15.6481 | -106.1074
15.6518 | -106.1011
15.6554 | -106.0952
15.6591 | -106.0894
15.6628 | -106.0836
15.6665 | -106.0742
15.6701 | -106.0624
15.6738 | -106.0506
15.6775 | -106.0383
15.6812 | -106.0274
15.6848 | -106.0174
15.6885 | -106.0089
15.6922 | -106
15.6959 | -105.9893
15.6995 | -105.9778
15.7032 | -105.9697
15.7069 | -105.9648
15.7105 | -105.9622
15.7142 | -105.9598
15.7179 | -105.9589
15.7216 | -105.9597
15.7252 | -105.9594
15.7289 | -105.9617
15.7326 | -105.9698
15.7363 | -105.9793
15.7399 | -105.9866
15.7436 | -105.9934
15.7473 | -106.0034
15.751 | -106.016
15.7546 | -106.0298
15.7583 | -106.0411
15.762 | -106.0564
15.7656 | -106.0705
15.7693 | -106.0836
15.773 | -106.096
15.7767 | -106.1063
15.7803 | -106.1144
15.784 | -106.1243
15.7877 | -106.1354
15.7914 | -106.1422
15.795 | -106.1494
15.7987 | -106.155
15.8024 | -106.159
15.8061 | -106.1608
15.8097 | -106.1618
15.8134 | -106.1583
15.8171 | -106.1558
15.8207 | -106.1506
15.8244 | -106.1475
15.8281 | -106.1395
15.8318 | -106.13
15.8354 | -106.1189
15.8391 | -106.1083
15.8428 | -106.0963
15.8465 | -106.0828
15.8501 | -106.0663
15.8538 | -106.049
15.8575 | -106.0316
15.8612 | -106.0134
15.8648 | -105.9975
15.8685 | -105.9843
15.8722 | -105.9708
15.8758 | -105.9557
15.8795 | -105.9413
15.8832 | -105.9282
15.8869 | -105.9151
15.8905 | -105.9045
15.8942 | -105.896
15.8979 | -105.8894
15.9016 | -105.8807
15.9052 | -105.8748
15.9089 | -105.869
15.9126 | -105.8645
15.9162 | -105.8603
15.9199 | -105.8588
15.9236 | -105.8593
15.9273 | -105.8616
15.9309 | -105.8614
15.9346 | -105.8641
15.9383 | -105.8686
15.942 | -105.8733
15.9456 | -105.8792
15.9493 | -105.8868
15.953 | -105.8975
15.9567 | -105.9087
15.9603 | -105.9202
15.964 | -105.9326
15.9677 | -105.9466
15.9713 | -105.963
15.975 | -105.9809
15.9787 | -105.9984
15.9824 | -106.015
15.986 | -106.0352
15.9897 | -106.0549
15.9934 | -106.0769
15.9971 | -106.1004
16.0007 | -106.1254
16.0044 | -106.1492
16.0081 | -106.1713
16.0118 | -106.1924
16.0154 | -106.2105
16.0191 | -106.2298
16.0228 | -106.248
16.0264 | -106.2638
16.0301 | -106.2772
16.0338 | -106.2919
16.0375 | -106.3049
16.0411 | -106.3166
16.0448 | -106.325
16.0485 | -106.3304
16.0522 | -106.3358
16.0558 | -106.3391
16.0595 | -106.3406
16.0632 | -106.3374
16.0669 | -106.3333
16.0705 | -106.3265
16.0742 | -106.3187
16.0779 | -106.3105
16.0815 | -106.3004
16.0852 | -106.2899
16.0889 | -106.2829
16.0926 | -106.2707
16.0962 | -106.2554
16.0999 | -106.239
16.1036 | -106.2223
16.1073 | -106.2045
16.1109 | -106.1876
16.1146 | -106.1729
16.1183 | -106.1566
16.122 | -106.1391
16.1256 | -106.1247
16.1293 | -106.1097
16.133 | -106.0936
16.1366 | -106.0763
16.1403 | -106.0575
16.144 | -106.0383
16.1477 | -106.0231
16.1513 | -106.0087
16.155 | -105.992
16.1587 | -105.9779
16.1624 | -105.9681
16.166 | -105.9595
16.1697 | -105.9531
16.1734 | -105.9479
16.1771 | -105.9419
16.1807 | -105.9373
16.1844 | -105.9373
16.1881 | -105.9406
16.1917 | -105.9453
16.1954 | -105.9539
16.1991 | -105.963
16.2028 | -105.9738
16.2064 | -105.9855
16.2101 | -105.998
16.2138 | -106.0112
16.2175 | -106.0276
16.2211 | -106.0452
16.2248 | -106.0647
16.2285 | -106.0829
16.2322 | -106.1031
16.2358 | -106.1267
16.2395 | -106.1516
16.2432 | -106.1748
16.2468 | -106.1981
16.2505 | -106.2203
16.2542 | -106.2437
16.2579 | -106.2682
16.2615 | -106.2921
16.2652 | -106.3148
16.2689 | -106.3354
16.2726 | -106.3564
16.2762 | -106.3759
16.2799 | -106.3959
16.2836 | -106.4155
16.2872 | -106.4336
16.2909 | -106.4504
16.2946 | -106.4683
16.2983 | -106.4821
16.3019 | -106.4954
16.3056 | -106.5082
16.3093 | -106.5193
16.313 | -106.529
16.3166 | -106.5368
16.3203 | -106.5416
16.324 | -106.5471
16.3277 | -106.55
16.3313 | -106.552
16.335 | -106.5544
16.3387 | -106.5583
16.3423 | -106.5594
16.346 | -106.5576
16.3497 | -106.5578
16.3534 | -106.5571
16.357 | -106.5539
16.3607 | -106.5521
16.3644 | -106.551
16.3681 | -106.5498
16.3717 | -106.5516
16.3754 | -106.5516
16.3791 | -106.5498
16.3828 | -106.5506
16.3864 | -106.5505
16.3901 | -106.5471
16.3938 | -106.5434
16.3974 | -106.5379
16.4011 | -106.5339
16.4048 | -106.5326
16.4085 | -106.5319
16.4121 | -106.5299
16.4158 | -106.527
16.4195 | -106.5254
16.4232 | -106.5231
16.4268 | -106.5223
16.4305 | -106.5202
16.4342 | -106.52
16.4379 | -106.5215
16.4415 | -106.5216
16.4452 | -106.5235
16.4489 | -106.5261
16.4525 | -106.5279
16.4562 | -106.5314
16.4599 | -106.5324
16.4636 | -106.5338
16.4672 | -106.5341
16.4709 | -106.5352
16.4746 | -106.5371
16.4783 | -106.5384
16.4819 | -106.5403
16.4856 | -106.5436
16.4893 | -106.5455
16.493 | -106.5477
16.4966 | -106.5504
16.5003 | -106.5552
16.504 | -106.5621
16.5076 | -106.567
16.5113 | -106.5677
16.515 | -106.5684
16.5187 | -106.568
16.5223 | -106.5709
16.526 | -106.5755
16.5297 | -106.5806
16.5334 | -106.5879
16.537 | -106.5945
16.5407 | -106.5987
16.5444 | -106.6012
16.5481 | -106.6072
16.5517 | -106.6144
16.5554 | -106.6214
16.5591 | -106.6277
16.5627 | -106.6356
16.5664 | -106.6411
16.5701 | -106.6446
16.5738 | -106.6458
16.5774 | -106.6493
16.5811 | -106.6529
16.5848 | -106.656
16.5885 | -106.6587
16.5921 | -106.6601
16.5958 | -106.6626
16.5995 | -106.6669
16.6031 | -106.669
16.6068 | -106.669
16.6105 | -106.6722
16.6142 | -106.6754
16.6178 | -106.6768
16.6215 | -106.6802
16.6252 | -106.6847
16.6289 | -106.6867
16.6325 | -106.6918
16.6362 | -106.6961
16.6399 | -106.7008
16.6436 | -106.7071
16.6472 | -106.7132
16.6509 | -106.7198
16.6546 | -106.7256
16.6582 | -106.7318
16.6619 | -106.7366
16.6656 | -106.7427
16.6693 | -106.7493
16.6729 | -106.7541
16.6766 | -106.7585
16.6803 | -106.7636
16.684 | -106.7698
16.6876 | -106.7787
16.6913 | -106.7872
16.695 | -106.7944
16.6987 | -106.8013
16.7023 | -106.8108
16.706 | -106.8203
16.7097 | -106.8288
16.7133 | -106.8355
16.717 | -106.8418
16.7207 | -106.8487
16.7244 | -106.8545
16.728 | -106.859
16.7317 | -106.8626
16.7354 | -106.8661
16.7391 | -106.8666
16.7427 | -106.8636
16.7464 | -106.8601
16.7501 | -106.8617
16.7538 | -106.862
16.7574 | -106.8582
16.7611 | -106.8495
16.7648 | -106.8366
16.7684 | -106.8232
16.7721 | -106.8118
16.7758 | -106.8007
16.7795 | -106.7869
16.7831 | -106.7753
16.7868 | -106.7616
16.7905 | -106.7456
16.7942 | -106.7277
16.7978 | -106.7102
16.8015 | -106.6951
16.8052 | -106.6824
16.8089 | -106.6683
16.8125 | -106.653
16.8162 | -106.6374
16.8199 | -106.6252
16.8235 | -106.6135
16.8272 | -106.6
16.8309 | -106.5874
16.8346 | -106.5754
16.8382 | -106.5638
16.8419 | -106.5564
16.8456 | -106.5509
16.8493 | -106.5442
16.8529 | -106.5362
16.8566 | -106.5311
16.8603 | -106.5302
16.864 | -106.5275
16.8676 | -106.5252
16.8713 | -106.5237
16.875 | -106.5242
16.8786 | -106.5241
16.8823 | -106.5237
16.886 | -106.5242
16.8897 | -106.5266
16.8933 | -106.5291
16.897 | -106.5318
16.9007 | -106.5354
16.9044 | -106.5386
16.908 | -106.5411
16.9117 | -106.5431
16.9154 | -106.5437
16.919 | -106.5435
16.9227 | -106.5438
16.9264 | -106.542
16.9301 | -106.5368
16.9337 | -106.5313
16.9374 | -106.5281
16.9411 | -106.5235
16.9448 | -106.518
16.9484 | -106.5118
16.9521 | -106.5039
16.9558 | -106.4944
16.9595 | -106.4903
16.9631 | -106.4855
16.9668 | -106.4793
16.9705 | -106.4749
16.9741 | -106.4703
16.9778 | -106.4637
16.9815 | -106.4589
16.9852 | -106.4592
16.9888 | -106.4607
16.9925 | -106.4578
16.9962 | -106.4559
16.9999 | -106.4574
17.0035 | -106.4614
17.0072 | -106.4677
17.0109 | -106.4729
17.0146 | -106.4763
17.0182 | -106.4816
17.0219 | -106.4912
17.0256 | -106.4983
17.0292 | -106.5048
17.0329 | -106.5157
17.0366 | -106.5314
17.0403 | -106.5467
17.0439 | -106.5602
17.0476 | -106.5746
17.0513 | -106.5908
17.055 | -106.6074
17.0586 | -106.6235
17.0623 | -106.6388
17.066 | -106.6549
17.0697 | -106.6736
17.0733 | -106.6899
17.077 | -106.7069
17.0807 | -106.7228
17.0843 | -106.7364
17.088 | -106.7481
17.0917 | -106.759
17.0954 | -106.7717
17.099 | -106.7824
17.1027 | -106.7903
17.1064 | -106.8004
17.1101 | -106.8095
17.1137 | -106.8139
17.1174 | -106.8169
17.1211 | -106.82
17.1248 | -106.819
17.1284 | -106.8174
17.1321 | -106.8137
17.1358 | -106.8097
17.1394 | -106.8075
17.1431 | -106.805
17.1468 | -106.7973
17.1505 | -106.7896
17.1541 | -106.7836
17.1578 | -106.7772
17.1615 | -106.7704
17.1652 | -106.7626
17.1688 | -106.7552
17.1725 | -106.7492
17.1762 | -106.7423
17.1799 | -106.7334
17.1835 | -106.7216
17.1872 | -106.7097
17.1909 | -106.6991
17.1945 | -106.6905
17.1982 | -106.681
17.2019 | -106.6691
17.2056 | -106.6588
17.2092 | -106.6466
17.2129 | -106.6334
17.2166 | -106.619
17.2203 | -106.6057
17.2239 | -106.5946
17.2276 | -106.5824
17.2313 | -106.5674
17.235 | -106.5541
17.2386 | -106.5439
17.2423 | -106.5334
17.246 | -106.5206
17.2496 | -106.5078
17.2533 | -106.4947
17.257 | -106.4845
17.2607 | -106.4737
17.2643 | -106.4619
17.268 | -106.4499
17.2717 | -106.4412
17.2754 | -106.4336
17.279 | -106.4244
17.2827 | -106.4131
17.2864 | -106.4031
17.2901 | -106.3945
17.2937 | -106.3885
17.2974 | -106.3825
17.3011 | -106.3774
17.3047 | -106.3734
17.3084 | -106.3688
17.3121 | -106.3649
17.3158 | -106.3619
17.3194 | -106.3599
17.3231 | -106.356
17.3268 | -106.3517
17.3305 | -106.3505
17.3341 | -106.3503
17.3378 | -106.3496
17.3415 | -106.3491
17.3451 | -106.3477
17.3488 | -106.3485
17.3525 | -106.3504
17.3562 | -106.3515
17.3598 | -106.354
17.3635 | -106.3562
17.3672 | -106.359
17.3709 | -106.3644
17.3745 | -106.3695
17.3782 | -106.3737
17.3819 | -106.378
17.3856 | -106.3814
17.3892 | -106.3849
17.3929 | -106.3911
17.3966 | -106.3972
17.4002 | -106.4015
17.4039 | -106.405
17.4076 | -106.4067
17.4113 | -106.4068
17.4149 | -106.4055
17.4186 | -106.4062
17.4223 | -106.4073
17.426 | -106.4085
17.4296 | -106.4116
17.4333 | -106.4125
17.437 | -106.4087
17.4407 | -106.4034
17.4443 | -106.397
17.448 | -106.3908
17.4517 | -106.3834
17.4553 | -106.3789
17.459 | -106.374
17.4627 | -106.3696
17.4664 | -106.3631
17.47 | -106.3537
17.4737 | -106.3445
17.4774 | -106.3366
17.4811 | -106.3259
17.4847 | -106.3111
17.4884 | -106.2985
17.4921 | -106.2836
17.4958 | -106.2713
17.4994 | -106.2609
17.5031 | -106.2483
17.5068 | -106.233
17.5104 | -106.2214
17.5141 | -106.2099
17.5178 | -106.1955
17.5215 | -106.1812
17.5251 | -106.1693
17.5288 | -106.1563
17.5325 | -106.1459
17.5362 | -106.1333
17.5398 | -106.1188
17.5435 | -106.1067
17.5472 | -106.0967
17.5509 | -106.0863
17.5545 | -106.0764
17.5582 | -106.0692
17.5619 | -106.0644
17.5655 | -106.0614
17.5692 | -106.0602
17.5729 | -106.0587
17.5766 | -106.0599
17.5802 | -106.0605
17.5839 | -106.0627
17.5876 | -106.0664
17.5913 | -106.0683
17.5949 | -106.0731
17.5986 | -106.0779
17.6023 | -106.0793
17.606 | -106.0855
17.6096 | -106.095
17.6133 | -106.1017
17.617 | -106.1067
17.6206 | -106.111
17.6243 | -106.1133
17.628 | -106.117
17.6317 | -106.1217
17.6353 | -106.1267
17.639 | -106.1298
17.6427 | -106.1312
17.6464 | -106.1329
17.65 | -106.1327
17.6537 | -106.1313
17.6574 | -106.1295
17.661 | -106.1269
17.6647 | -106.1227
17.6684 | -106.1145
17.6721 | -106.1045
17.6757 | -106.0933
17.6794 | -106.082
17.6831 | -106.0714
17.6868 | -106.0562
17.6904 | -106.0394
17.6941 | -106.0208
17.6978 | -105.9976
17.7015 | -105.9746
17.7051 | -105.9479
17.7088 | -105.9186
17.7125 | -105.8901
17.7161 | -105.8612
17.7198 | -105.8297
17.7235 | -105.7951
17.7272 | -105.7629
17.7308 | -105.7302
17.7345 | -105.6966
17.7382 | -105.6605
17.7419 | -105.6261
17.7455 | -105.5912
17.7492 | -105.5528
17.7529 | -105.5154
17.7566 | -105.4799
17.7602 | -105.4471
17.7639 | -105.4127
17.7676 | -105.378
17.7712 | -105.3462
17.7749 | -105.3139
17.7786 | -105.2819
17.7823 | -105.2518
17.7859 | -105.222
17.7896 | -105.1958
17.7933 | -105.1717
17.797 | -105.1478
17.8006 | -105.1259
17.8043 | -105.1059
17.808 | -105.0884
17.8117 | -105.0713
17.8153 | -105.0537
17.819 | -105.0374
17.8227 | -105.0237
17.8263 | -105.0136
17.83 | -105.0068
17.8337 | -105.0036
17.8374 | -105.0002
17.841 | -104.9993
17.8447 | -105.0003
17.8484 | -105.0007
17.8521 | -105.0045
17.8557 | -105.013
17.8594 | -105.0198
17.8631 | -105.0258
17.8668 | -105.0365
17.8704 | -105.0498
17.8741 | -105.0638
17.8778 | -105.081
17.8814 | -105.0961
17.8851 | -105.1095
17.8888 | -105.1245
17.8925 | -105.1401
17.8961 | -105.1529
17.8998 | -105.1648
17.9035 | -105.1771
17.9072 | -105.1894
17.9108 | -105.2008
17.9145 | -105.2108
17.9182 | -105.2217
17.9219 | -105.2316
17.9255 | -105.2401
17.9292 | -105.246
17.9329 | -105.247
17.9365 | -105.2452
17.9402 | -105.2457
17.9439 | -105.2452
17.9476 | -105.243
17.9512 | -105.2402
17.9549 | -105.2297
17.9586 | -105.2132
17.9623 | -105.1932
17.9659 | -105.1702
17.9696 | -105.1476
17.9733 | -105.1224
17.9769 | -105.0973
17.9806 | -105.0686
17.9843 | -105.0392
17.988 | -105.0104
17.9916 | -104.9809
17.9953 | -104.9488
17.999 | -104.9159
18.0027 | -104.8824
18.0063 | -104.8433
18.01 | -104.8051
18.0137 | -104.7676
18.0174 | -104.7276
18.021 | -104.6908
18.0247 | -104.6548
18.0284 | -104.6178
18.032 | -104.5823
18.0357 | -104.5472
18.0394 | -104.5134
18.0431 | -104.479
18.0467 | -104.4476
18.0504 | -104.4173
18.0541 | -104.386
18.0578 | -104.3557
18.0614 | -104.3284
18.0651 | -104.302
18.0688 | -104.2771
18.0725 | -104.2559
18.0761 | -104.2393
18.0798 | -104.2241
18.0835 | -104.2107
18.0871 | -104.199
18.0908 | -104.1862
18.0945 | -104.1793
18.0982 | -104.1753
18.1018 | -104.1727
18.1055 | -104.1728
18.1092 | -104.1722
18.1129 | -104.1745
18.1165 | -104.1798
18.1202 | -104.1874
18.1239 | -104.1939
18.1276 | -104.201
18.1312 | -104.2091
18.1349 | -104.2163
18.1386 | -104.2256
18.1422 | -104.2349
18.1459 | -104.2429
18.1496 | -104.251
18.1533 | -104.2606
18.1569 | -104.2701
18.1606 | -104.2787
18.1643 | -104.2832
18.168 | -104.2864
18.1716 | -104.2904
18.1753 | -104.2934
18.179 | -104.2955
18.1827 | -104.2971
18.1863 | -104.2988
18.19 | -104.3025
18.1937 | -104.3055
18.1973 | -104.3063
18.201 | -104.3053
18.2047 | -104.3055
18.2084 | -104.3065
18.212 | -104.3101
18.2157 | -104.3157
18.2194 | -104.3231
18.2231 | -104.3278
18.2267 | -104.3292
18.2304 | -104.3346
18.2341 | -104.3391
18.2378 | -104.341
18.2414 | -104.3449
18.2451 | -104.3489
18.2488 | -104.3524
18.2524 | -104.3577
18.2561 | -104.3638
18.2598 | -104.3702
18.2635 | -104.378
18.2671 | -104.3841
18.2708 | -104.3873
18.2745 | -104.3905
18.2782 | -104.3959
18.2818 | -104.4012
18.2855 | -104.4057
18.2892 | -104.4112
18.2929 | -104.4153
18.2965 | -104.4176
18.3002 | -104.4165
18.3039 | -104.4157
18.3075 | -104.4129
18.3112 | -104.4113
18.3149 | -104.4083
18.3186 | -104.403
18.3222 | -104.3942
18.3259 | -104.3855
18.3296 | -104.3773
18.3333 | -104.3684
18.3369 | -104.357
18.3406 | -104.3439
18.3443 | -104.3299
18.348 | -104.3181
18.3516 | -104.308
18.3553 | -104.2959
18.359 | -104.2836
18.3626 | -104.27
18.3663 | -104.2542
18.37 | -104.242
18.3737 | -104.2295
18.3773 | -104.2188
18.381 | -104.2081
18.3847 | -104.1985
18.3884 | -104.1893
18.392 | -104.181
18.3957 | -104.1737
18.3994 | -104.1665
18.403 | -104.1611
18.4067 | -104.158
18.4104 | -104.1562
18.4141 | -104.1551
18.4177 | -104.1531
18.4214 | -104.1525
18.4251 | -104.1548
18.4288 | -104.1584
18.4324 | -104.1598
18.4361 | -104.1593
18.4398 | -104.161
18.4435 | -104.1654
18.4471 | -104.1703
18.4508 | -104.1754
18.4545 | -104.1847
18.4581 | -104.1953
18.4618 | -104.2047
18.4655 | -104.214
18.4692 | -104.223
18.4728 | -104.2328
18.4765 | -104.2442
18.4802 | -104.2549
18.4839 | -104.2661
18.4875 | -104.2788
18.4912 | -104.2902
18.4949 | -104.3005
18.4986 | -104.3142
18.5022 | -104.328
18.5059 | -104.3406
18.5096 | -104.353
18.5132 | -104.3651
18.5169 | -104.3782
18.5206 | -104.3913
18.5243 | -104.4044
18.5279 | -104.4166
18.5316 | -104.4298
18.5353 | -104.4439
18.539 | -104.4539
18.5426 | -104.4627
18.5463 | -104.4763
18.55 | -104.487
18.5537 | -104.4964
18.5573 | -104.5046
18.561 | -104.5132
18.5647 | -104.5257
18.5683 | -104.5372
18.572 | -104.5487
18.5757 | -104.5594
18.5794 | -104.5707
18.583 | -104.5825
18.5867 | -104.5941
18.5904 | -104.6042
18.5941 | -104.6142
18.5977 | -104.6217
18.6014 | -104.6304
18.6051 | -104.6411
18.6088 | -104.652
18.6124 | -104.6607
18.6161 | -104.6695
18.6198 | -104.6765
18.6234 | -104.6813
18.6271 | -104.6905
18.6308 | -104.7014
18.6345 | -104.7081
18.6381 | -104.7094
18.6418 | -104.7096
18.6455 | -104.711
18.6492 | -104.7136
18.6528 | -104.7167
18.6565 | -104.7159
18.6602 | -104.7161
18.6639 | -104.7167
18.6675 | -104.7176
18.6712 | -104.7146
18.6749 | -104.7105
18.6785 | -104.7061
18.6822 | -104.7048
18.6859 | -104.7035
18.6896 | -104.7015
18.6932 | -104.6976
18.6969 | -104.6943
18.7006 | -104.6943
18.7043 | -104.692
18.7079 | -104.6868
18.7116 | -104.6808
18.7153 | -104.6756
18.7189 | -104.665
18.7226 | -104.6557
18.7263 | -104.6511
18.73 | -104.6465
18.7336 | -104.6416
18.7373 | -104.6355
18.741 | -104.6293
18.7447 | -104.6223
18.7483 | -104.6152
18.752 | -104.6084
18.7557 | -104.5996
18.7594 | -104.593
18.763 | -104.5849
18.7667 | -104.5771
18.7704 | -104.5712
18.774 | -104.5666
18.7777 | -104.5613
18.7814 | -104.5565
18.7851 | -104.5501
18.7887 | -104.5423
18.7924 | -104.538
18.7961 | -104.5363
18.7998 | -104.5343
18.8034 | -104.5342
18.8071 | -104.5345
18.8108 | -104.5317
18.8145 | -104.53
18.8181 | -104.5298
18.8218 | -104.5306
18.8255 | -104.5336
18.8291 | -104.5368
18.8328 | -104.5377
18.8365 | -104.5393
18.8402 | -104.539
18.8438 | -104.5393
18.8475 | -104.5419
18.8512 | -104.5447
18.8549 | -104.5468
18.8585 | -104.5478
18.8622 | -104.5492
18.8659 | -104.5524
18.8696 | -104.5531
18.8732 | -104.5538
18.8769 | -104.555
18.8806 | -104.5541
18.8842 | -104.5533
18.8879 | -104.5545
18.8916 | -104.5541
18.8953 | -104.5524
18.8989 | -104.5517
18.9026 | -104.5529
18.9063 | -104.5565
18.91 | -104.5567
18.9136 | -104.5546
18.9173 | -104.5502
18.921 | -104.5447
18.9247 | -104.5424
18.9283 | -104.5365
18.932 | -104.5287
18.9357 | -104.5208
18.9393 | -104.5143
18.943 | -104.5075
18.9467 | -104.4985
18.9504 | -104.485
18.954 | -104.4717
18.9577 | -104.4562
18.9614 | -104.4391
18.9651 | -104.4217
18.9687 | -104.4055
18.9724 | -104.3877
18.9761 | -104.3682
18.9798 | -104.3494
18.9834 | -104.3283
18.9871 | -104.3095
18.9908 | -104.2884
18.9944 | -104.2637
18.9981 | -104.2365
19.0018 | -104.2096
19.0055 | -104.1839
19.0091 | -104.1582
19.0128 | -104.1335
19.0165 | -104.1092
19.0202 | -104.0837
19.0238 | -104.0595
19.0275 | -104.0337
19.0312 | -104.0072
19.0348 | -103.98
19.0385 | -103.951
19.0422 | -103.9206
19.0459 | -103.8914
19.0495 | -103.8654
19.0532 | -103.8442
19.0569 | -103.8222
19.0606 | -103.7978
19.0642 | -103.7731
19.0679 | -103.7495
19.0716 | -103.7271
19.0753 | -103.7108
19.0789 | -103.6866
19.0826 | -103.6639
19.0863 | -103.6443
19.0899 | -103.6269
19.0936 | -103.6105
19.0973 | -103.5883
19.101 | -103.5669
19.1046 | -103.5467
19.1083 | -103.5272
19.112 | -103.5105
19.1157 | -103.4932
19.1193 | -103.4754
19.123 | -103.4583
19.1267 | -103.4419
19.1304 | -103.4261
19.134 | -103.4084
19.1377 | -103.3891
19.1414 | -103.3692
19.145 | -103.3505
19.1487 | -103.3331
19.1524 | -103.3144
19.1561 | -103.297
19.1597 | -103.2796
19.1634 | -103.2637
19.1671 | -103.2472
19.1708 | -103.2292
19.1744 | -103.2092
19.1781 | -103.1914
19.1818 | -103.1755
19.1855 | -103.1609
19.1891 | -103.1474
19.1928 | -103.1358
19.1965 | -103.1245
19.2001 | -103.1141
19.2038 | -103.1043
19.2075 | -103.0956
19.2112 | -103.0866
19.2148 | -103.0785
19.2185 | -103.0739
19.2222 | -103.0702
19.2259 | -103.0663
19.2295 | -103.0662
19.2332 | -103.0684
19.2369 | -103.0705
19.2406 | -103.0709
19.2442 | -103.0751
19.2479 | -103.081
19.2516 | -103.0873
19.2552 | -103.0945
19.2589 | -103.1021
19.2626 | -103.1109
19.2663 | -103.1211
19.2699 | -103.1342
19.2736 | -103.1498
19.2773 | -103.1646
19.281 | -103.178
19.2846 | -103.1924
19.2883 | -103.2087
19.292 | -103.2263
19.2957 | -103.2443
19.2993 | -103.2623
19.303 | -103.2802
19.3067 | -103.2994
19.3103 | -103.3194
19.314 | -103.3391
19.3177 | -103.358
19.3214 | -103.3726
19.325 | -103.3841
19.3287 | -103.3998
19.3324 | -103.4138
19.3361 | -103.4247
19.3397 | -103.4353
19.3434 | -103.4468
19.3471 | -103.457
19.3508 | -103.4668
19.3544 | -103.4776
19.3581 | -103.4878
19.3618 | -103.4935
19.3654 | -103.4956
19.3691 | -103.4954
19.3728 | -103.4984
19.3765 | -103.499
19.3801 | -103.4996
19.3838 | -103.4991
19.3875 | -103.4943
19.3912 | -103.4878
19.3948 | -103.4812
19.3985 | -103.4748
19.4022 | -103.4691
19.4059 | -103.463
19.4095 | -103.4535
19.4132 | -103.4465
19.4169 | -103.4429
19.4205 | -103.4332
19.4242 | -103.4232
19.4279 | -103.4155
19.4316 | -103.4067
19.4352 | -103.3973
19.4389 | -103.387
19.4426 | -103.3757
19.4463 | -103.3666
19.4499 | -103.3579
19.4536 | -103.3476
19.4573 | -103.3374
19.4609 | -103.3251
19.4646 | -103.311
19.4683 | -103.2944
19.472 | -103.2785
19.4756 | -103.2644
19.4793 | -103.2487
19.483 | -103.232
19.4867 | -103.2192
19.4903 | -103.2072
19.494 | -103.1913
19.4977 | -103.1729
19.5014 | -103.1543
19.505 | -103.1344
19.5087 | -103.1123
19.5124 | -103.093
19.516 | -103.0767
19.5197 | -103.0579
19.5234 | -103.0405
19.5271 | -103.022
19.5307 | -103.0031
19.5344 | -102.9844
19.5381 | -102.9656
19.5418 | -102.948
19.5454 | -102.9289
19.5491 | -102.9097
19.5528 | -102.8942
19.5565 | -102.8797
19.5601 | -102.8662
19.5638 | -102.8532
19.5675 | -102.8413
19.5711 | -102.8308
19.5748 | -102.8225
19.5785 | -102.8153
19.5822 | -102.8075
19.5858 | -102.8013
19.5895 | -102.7972
19.5932 | -102.7953
19.5969 | -102.795
19.6005 | -102.7979
19.6042 | -102.8127
19.6079 | -102.815
19.6116 | -102.8199
19.6152 | -102.8282
19.6189 | -102.838
19.6226 | -102.8482
19.6262 | -102.8585
19.6299 | -102.8718
19.6336 | -102.8883
19.6373 | -102.905
19.6409 | -102.9237
19.6446 | -102.9431
19.6483 | -102.9652
19.652 | -102.9913
19.6556 | -103.0181
19.6593 | -103.0423
19.663 | -103.0683
19.6667 | -103.0964
19.6703 | -103.1236
19.674 | -103.1485
19.6777 | -103.1723
19.6813 | -103.196
19.685 | -103.2201
19.6887 | -103.2457
19.6924 | -103.2697
19.696 | -103.289
19.6997 | -103.3092
19.7034 | -103.3274
19.7071 | -103.3463
19.7107 | -103.3619
19.7144 | -103.3764
19.7181 | -103.3892
19.7218 | -103.3993
19.7254 | -103.4057
19.7291 | -103.4115
19.7328 | -103.4168
19.7364 | -103.4217
19.7401 | -103.423
19.7438 | -103.4236
19.7475 | -103.4255
19.7511 | -103.4245
19.7548 | -103.421
19.7585 | -103.4184
19.7622 | -103.4143
19.7658 | -103.4057
19.7695 | -103.3946
19.7732 | -103.3837
19.7768 | -103.377
19.7805 | -103.3708
19.7842 | -103.3608
19.7879 | -103.3533
19.7915 | -103.3439
19.7952 | -103.3328
19.7989 | -103.324
19.8026 | -103.3145
19.8062 | -103.3063
19.8099 | -103.2991
19.8136 | -103.2903
19.8173 | -103.2844
19.8209 | -103.2766
19.8246 | -103.2683
19.8283 | -103.2608
19.8319 | -103.2511
19.8356 | -103.2428
19.8393 | -103.2356
19.843 | -103.2275
19.8466 | -103.221
19.8503 | -103.2143
19.854 | -103.2089
19.8577 | -103.2057
19.8613 | -103.2038
19.865 | -103.202
19.8687 | -103.2
19.8724 | -103.1993
19.876 | -103.2003
19.8797 | -103.2016
19.8834 | -103.2029
19.887 | -103.206
19.8907 | -103.2111
19.8944 | -103.2171
19.8981 | -103.2217
19.9017 | -103.2279
19.9054 | -103.2345
19.9091 | -103.2441
19.9128 | -103.2504
19.9164 | -103.2566
19.9201 | -103.2662
19.9238 | -103.2763
19.9275 | -103.2885
19.9311 | -103.2997
19.9348 | -103.3096
19.9385 | -103.3223
19.9421 | -103.3359
19.9458 | -103.345
19.9495 | -103.3563
19.9532 | -103.3707
19.9568 | -103.3871
19.9605 | -103.4037
19.9642 | -103.4188
19.9679 | -103.4311
19.9715 | -103.4434
19.9752 | -103.4528
19.9789 | -103.4613
19.9826 | -103.4721
19.9862 | -103.4807
19.9899 | -103.4867
19.9936 | -103.4933
19.9972 | -103.4982
20.0009 | -103.5042
20.0046 | -103.5095
20.0083 | -103.515
20.0119 | -103.5212
20.0156 | -103.5241
20.0193 | -103.5236
20.023 | -103.5263
20.0266 | -103.5291
20.0303 | -103.5295
20.034 | -103.5314
20.0377 | -103.5313
20.0413 | -103.5294
20.045 | -103.5269
20.0487 | -103.5251
20.0523 | -103.5253
20.056 | -103.5254
20.0597 | -103.5228
20.0634 | -103.5191
20.067 | -103.5187
20.0707 | -103.5162
20.0744 | -103.5117
20.0781 | -103.5076
20.0817 | -103.5034
20.0854 | -103.4994
20.0891 | -103.4964
20.0927 | -103.4933
20.0964 | -103.4907
20.1001 | -103.4871
20.1038 | -103.4852
20.1074 | -103.4843
20.1111 | -103.4813
20.1148 | -103.4776
20.1185 | -103.4728
20.1221 | -103.4686
20.1258 | -103.4661
20.1295 | -103.464
20.1332 | -103.4625
20.1368 | -103.4605
20.1405 | -103.4577
20.1442 | -103.4534
20.1478 | -103.4495
20.1515 | -103.4483
20.1552 | -103.4476
20.1589 | -103.4454
20.1625 | -103.4415
20.1662 | -103.4376
20.1699 | -103.4363
20.1736 | -103.4368
20.1772 | -103.4356
20.1809 | -103.4338
20.1846 | -103.4319
20.1883 | -103.4287
20.1919 | -103.4271
20.1956 | -103.4263
20.1993 | -103.4281
20.2029 | -103.4293
20.2066 | -103.4292
20.2103 | -103.43
20.214 | -103.4306
20.2176 | -103.4314
20.2213 | -103.4309
20.225 | -103.4282
20.2287 | -103.4278
20.2323 | -103.4312
20.236 | -103.4312
20.2397 | -103.4298
20.2434 | -103.4278
20.247 | -103.4239
20.2507 | -103.4202
20.2544 | -103.4179
20.258 | -103.4161
20.2617 | -103.4161
20.2654 | -103.4166
20.2691 | -103.4156
20.2727 | -103.4152
20.2764 | -103.4117
20.2801 | -103.411
20.2838 | -103.4103
20.2874 | -103.4076
20.2911 | -103.4055
20.2948 | -103.4027
20.2985 | -103.3979
20.3021 | -103.394
20.3058 | -103.3914
20.3095 | -103.389
20.3131 | -103.3873
20.3168 | -103.3856
20.3205 | -103.3814
20.3242 | -103.3772
20.3278 | -103.3731
20.3315 | -103.3716
20.3352 | -103.3699
20.3389 | -103.3683
20.3425 | -103.3661
20.3462 | -103.3623
20.3499 | -103.3584
20.3536 | -103.3527
20.3572 | -103.3479
20.3609 | -103.344
20.3646 | -103.3384
20.3682 | -103.3307
20.3719 | -103.3228
20.3756 | -103.3155
20.3793 | -103.3093
20.3829 | -103.3032
20.3866 | -103.2953
20.3903 | -103.2863
20.394 | -103.2768
20.3976 | -103.2684
20.4013 | -103.2612
20.405 | -103.2533
20.4087 | -103.2383
20.4123 | -103.2236
20.416 | -103.2097
20.4197 | -103.1967
20.4233 | -103.1825
20.427 | -103.1702
20.4307 | -103.1543
20.4344 | -103.1373
20.438 | -103.1207
20.4417 | -103.1027
20.4454 | -103.0831
20.4491 | -103.0631
20.4527 | -103.0447
20.4564 | -103.0259
20.4601 | -103.0058
20.4637 | -102.9857
20.4674 | -102.966
20.4711 | -102.9462
20.4748 | -102.9274
20.4784 | -102.9076
20.4821 | -102.8856
20.4858 | -102.8638
20.4895 | -102.8421
20.4931 | -102.8221
20.4968 | -102.8006
20.5005 | -102.78
20.5042 | -102.7608
20.5078 | -102.7406
20.5115 | -102.7224
20.5152 | -102.7052
20.5188 | -102.6863
20.5225 | -102.6647
20.5262 | -102.6453
20.5299 | -102.6272
20.5335 | -102.612
20.5372 | -102.5995
20.5409 | -102.5861
20.5446 | -102.572
20.5482 | -102.5576
20.5519 | -102.5414
20.5556 | -102.527
20.5593 | -102.5121
20.5629 | -102.4995
20.5666 | -102.4899
20.5703 | -102.4794
20.5739 | -102.4685
20.5776 | -102.4578
20.5813 | -102.4498
20.585 | -102.44
20.5886 | -102.4309
20.5923 | -102.424
20.596 | -102.4173
20.5997 | -102.4095
20.6033 | -102.4002
20.607 | -102.392
20.6107 | -102.3856
20.6144 | -102.382
20.618 | -102.3779
20.6217 | -102.371
20.6254 | -102.3608
20.629 | -102.3461
20.6327 | -102.3344
20.6364 | -102.3226
20.6401 | -102.3069
20.6437 | -102.293
20.6474 | -102.2767
20.6511 | -102.259
20.6548 | -102.2434
20.6584 | -102.2301
20.6621 | -102.2148
20.6658 | -102.1885
20.6695 | -102.1677
20.6731 | -102.1454
20.6768 | -102.1242
20.6805 | -102.1031
20.6841 | -102.0807
20.6878 | -102.0578
20.6915 | -102.0347
20.6952 | -102.0116
20.6988 | -101.9878
20.7025 | -101.9651
20.7062 | -101.9424
20.7099 | -101.9184
20.7135 | -101.8936
20.7172 | -101.8684
20.7209 | -101.8414
20.7246 | -101.8161
20.7282 | -101.7893
20.7319 | -101.7634
20.7356 | -101.7392
20.7392 | -101.7154
20.7429 | -101.691
20.7466 | -101.6648
20.7503 | -101.6396
20.7539 | -101.6151
20.7576 | -101.5897
20.7613 | -101.5648
20.765 | -101.5398
20.7686 | -101.5158
20.7723 | -101.4955
20.776 | -101.4758
20.7797 | -101.4555
20.7833 | -101.4367
20.787 | -101.4172
20.7907 | -101.3989
20.7943 | -101.3811
20.798 | -101.3629
20.8017 | -101.3457
20.8054 | -101.331
20.809 | -101.3151
20.8127 | -101.3019
20.8164 | -101.2908
20.8201 | -101.2802
20.8237 | -101.2697
20.8274 | -101.2595
20.8311 | -101.2485
20.8347 | -101.2396
20.8384 | -101.2329
20.8421 | -101.2263
20.8458 | -101.2208
20.8494 | -101.2145
20.8531 | -101.2078
20.8568 | -101.2002
20.8605 | -101.195
20.8641 | -101.1905
20.8678 | -101.1888
20.8715 | -101.1886
20.8752 | -101.1848
20.8788 | -101.1835
20.8825 | -101.1861
20.8862 | -101.1883
20.8898 | -101.1916
20.8935 | -101.1966
20.8972 | -101.2037
20.9009 | -101.2108
20.9045 | -101.2182
20.9082 | -101.2269
20.9119 | -101.2369
20.9156 | -101.2476
20.9192 | -101.2615
20.9229 | -101.2786
20.9266 | -101.2949
20.9303 | -101.3117
20.9339 | -101.3308
20.9376 | -101.3507
20.9413 | -101.372
20.9449 | -101.3934
20.9486 | -101.4143
20.9523 | -101.4362
20.956 | -101.4596
20.9596 | -101.4832
20.9633 | -101.5101
20.967 | -101.5379
20.9707 | -101.5651
20.9743 | -101.5912
20.978 | -101.6184
20.9817 | -101.6474
20.9854 | -101.6788
20.989 | -101.7093
20.9927 | -101.738
20.9964 | -101.7675
21 | -101.7985
21.0037 | -101.8272
21.0074 | -101.856
21.0111 | -101.8868
21.0147 | -101.919
21.0184 | -101.9499
21.0221 | -101.9802
21.0258 | -102.0136
21.0294 | -102.0442
21.0331 | -102.0723
21.0368 | -102.1007
21.0405 | -102.1294
21.0441 | -102.159
21.0478 | -102.1871
21.0515 | -102.2144
21.0551 | -102.2413
21.0588 | -102.2666
21.0625 | -102.2904
21.0662 | -102.3131
21.0698 | -102.333
21.0735 | -102.3508
21.0772 | -102.3671
21.0809 | -102.3829
21.0845 | -102.3995
21.0882 | -102.4145
21.0919 | -102.4272
21.0956 | -102.4402
21.0992 | -102.4534
21.1029 | -102.4618
21.1066 | -102.4718
21.1102 | -102.4799
21.1139 | -102.4862
21.1176 | -102.4919
21.1213 | -102.4985
21.1249 | -102.5025
21.1286 | -102.5083
21.1323 | -102.5135
21.136 | -102.5176
21.1396 | -102.5234
21.1433 | -102.5296
21.147 | -102.5372
21.1506 | -102.5466
21.1543 | -102.5572
21.158 | -102.5659
21.1617 | -102.5749
21.1653 | -102.5826
21.169 | -102.5915
21.1727 | -102.6042
21.1764 | -102.6167
21.18 | -102.625
21.1837 | -102.6365
21.1874 | -102.6494
21.1911 | -102.6621
21.1947 | -102.6763
21.1984 | -102.688
21.2021 | -102.7018
21.2057 | -102.7172
21.2094 | -102.7314
21.2131 | -102.747
21.2168 | -102.7638
21.2204 | -102.7809
21.2241 | -102.7981
21.2278 | -102.8149
21.2315 | -102.8324
21.2351 | -102.8488
21.2388 | -102.8616
21.2425 | -102.8728
21.2462 | -102.8868
21.2498 | -102.8993
21.2535 | -102.9078
21.2572 | -102.9173
21.2608 | -102.928
21.2645 | -102.9368
21.2682 | -102.9449
21.2719 | -102.9531
21.2755 | -102.9589
21.2792 | -102.9627
21.2829 | -102.9639
21.2866 | -102.9629
21.2902 | -102.9636
21.2939 | -102.9652
21.2976 | -102.9657
21.3013 | -102.9633
21.3049 | -102.9601
21.3086 | -102.9541
21.3123 | -102.9473
21.3159 | -102.9405
21.3196 | -102.9351
21.3233 | -102.9286
21.327 | -102.9187
21.3306 | -102.9075
21.3343 | -102.8966
21.338 | -102.8848
21.3417 | -102.8695
21.3453 | -102.8554
21.349 | -102.8418
21.3527 | -102.8283
21.3564 | -102.816
21.36 | -102.8043
21.3637 | -102.7924
21.3674 | -102.7779
21.371 | -102.7639
21.3747 | -102.7516
21.3784 | -102.7395
21.3821 | -102.7256
21.3857 | -102.7098
21.3894 | -102.6964
21.3931 | -102.6829
21.3968 | -102.6693
21.4004 | -102.657
21.4041 | -102.645
21.4078 | -102.6313
21.4115 | -102.619
21.4151 | -102.6084
21.4188 | -102.5953
21.4225 | -102.5812
21.4261 | -102.569
21.4298 | -102.5567
21.4335 | -102.5451
21.4372 | -102.5365
21.4408 | -102.5292
21.4445 | -102.5211
21.4482 | -102.5114
21.4519 | -102.5041
21.4555 | -102.4964
21.4592 | -102.4887
21.4629 | -102.4815
21.4666 | -102.4751
21.4702 | -102.4694
21.4739 | -102.4656
21.4776 | -102.4621
21.4812 | -102.4596
21.4849 | -102.4572
21.4886 | -102.4544
21.4923 | -102.4528
21.4959 | -102.4537
21.4996 | -102.4552
21.5033 | -102.4543
21.507 | -102.4533
21.5106 | -102.4539
21.5143 | -102.4563
21.518 | -102.4585
21.5216 | -102.4617
21.5253 | -102.4647
21.529 | -102.4693
21.5327 | -102.4743
21.5363 | -102.481
21.54 | -102.4899
21.5437 | -102.4984
21.5474 | -102.509
21.551 | -102.5168
21.5547 | -102.5243
21.5584 | -102.535
21.5621 | -102.5464
21.5657 | -102.5594
21.5694 | -102.5728
21.5731 | -102.5861
21.5767 | -102.6027
21.5804 | -102.621
21.5841 | -102.6401
21.5878 | -102.66
21.5914 | -102.6786
21.5951 | -102.6958
21.5988 | -102.7144
21.6025 | -102.7346
21.6061 | -102.7554
21.6098 | -102.7746
21.6135 | -102.797
21.6172 | -102.8173
21.6208 | -102.8361
21.6245 | -102.8527
21.6282 | -102.87
21.6318 | -102.8875
21.6355 | -102.9024
21.6392 | -102.9175
21.6429 | -102.9333
21.6465 | -102.9523
21.6502 | -102.9698
21.6539 | -102.9849
21.6576 | -102.9972
21.6612 | -103.0084
21.6649 | -103.0204
21.6686 | -103.0327
21.6723 | -103.0443
21.6759 | -103.053
21.6796 | -103.0615
21.6833 | -103.0694
21.6869 | -103.0772
21.6906 | -103.0861
21.6943 | -103.0927
21.698 | -103.0992
21.7016 | -103.1032
21.7053 | -103.106
21.709 | -103.1115
21.7127 | -103.1171
21.7163 | -103.121
21.72 | -103.1251
21.7237 | -103.1299
21.7274 | -103.1309
21.731 | -103.1358
21.7347 | -103.1405
21.7384 | -103.1459
21.742 | -103.1507
21.7457 | -103.1506
21.7494 | -103.1504
21.7531 | -103.1528
21.7567 | -103.1578
21.7604 | -103.1612
21.7641 | -103.1649
21.7678 | -103.1705
21.7714 | -103.1778
21.7751 | -103.1869
21.7788 | -103.1959
21.7825 | -103.2041
21.7861 | -103.2129
21.7898 | -103.2231
21.7935 | -103.2329
21.7971 | -103.243
21.8008 | -103.2532
21.8045 | -103.2613
21.8082 | -103.2715
21.8118 | -103.2841
21.8155 | -103.2964
21.8192 | -103.3078
21.8229 | -103.3174
21.8265 | -103.3277
21.8302 | -103.3384
21.8339 | -103.3507
21.8376 | -103.3619
21.8412 | -103.3726
21.8449 | -103.3823
21.8486 | -103.3923
21.8522 | -103.405
21.8559 | -103.4147
21.8596 | -103.4241
21.8633 | -103.4348
21.8669 | -103.4437
21.8706 | -103.4523
21.8743 | -103.4593
21.878 | -103.4679
21.8816 | -103.4765
21.8853 | -103.4835
21.889 | -103.4912
21.8926 | -103.4994
21.8963 | -103.5092
21.9 | -103.52
21.9037 | -103.5324
21.9073 | -103.5443
21.911 | -103.5551
21.9147 | -103.5662
21.9184 | -103.5777
21.922 | -103.5879
21.9257 | -103.5997
21.9294 | -103.6124
21.9331 | -103.6252
21.9367 | -103.6376
21.9404 | -103.648
21.9441 | -103.66
21.9477 | -103.6683
21.9514 | -103.6762
21.9551 | -103.6815
21.9588 | -103.6833
21.9624 | -103.6851
21.9661 | -103.6885
21.9698 | -103.6893
21.9735 | -103.6911
21.9771 | -103.6935
21.9808 | -103.6959
21.9845 | -103.6969
21.9882 | -103.6983
21.9918 | -103.6983
21.9955 | -103.6966
21.9992 | -103.6939
22.0028 | -103.6918
22.0065 | -103.6899
22.0102 | -103.6879
22.0139 | -103.6827
22.0175 | -103.676
22.0212 | -103.6701
22.0249 | -103.6622
22.0286 | -103.6546
22.0322 | -103.6468
22.0359 | -103.6398
22.0396 | -103.6324
22.0433 | -103.6245
22.0469 | -103.6162
22.0506 | -103.6064
22.0543 | -103.5974
22.0579 | -103.5913
22.0616 | -103.5844
22.0653 | -103.5782
22.069 | -103.5705
22.0726 | -103.5609
22.0763 | -103.5516
22.08 | -103.5414
22.0837 | -103.5321
22.0873 | -103.5234
22.091 | -103.5156
22.0947 | -103.5068
22.0984 | -103.4999
22.102 | -103.4912
22.1057 | -103.485
22.1094 | -103.4807
22.113 | -103.4782
22.1167 | -103.4784
22.1204 | -103.4776
22.1241 | -103.4752
22.1277 | -103.4739
22.1314 | -103.4734
22.1351 | -103.4759
22.1388 | -103.4792
22.1424 | -103.4835
22.1461 | -103.4881
22.1498 | -103.4924
22.1535 | -103.4973
22.1571 | -103.5039
22.1608 | -103.5104
22.1645 | -103.5181
22.1681 | -103.5272
22.1718 | -103.537
22.1755 | -103.545
22.1792 | -103.5535
22.1828 | -103.5631
22.1865 | -103.5719
22.1902 | -103.5802
22.1939 | -103.5892
22.1975 | -103.5982
22.2012 | -103.6079
22.2049 | -103.6171
22.2085 | -103.6269
22.2122 | -103.6379
22.2159 | -103.6492
22.2196 | -103.6614
22.2232 | -103.673
22.2269 | -103.683
22.2306 | -103.6962
22.2343 | -103.7084
22.2379 | -103.72
22.2416 | -103.7305
22.2453 | -103.7431
22.249 | -103.7565
22.2526 | -103.7703
22.2563 | -103.7815
22.26 | -103.7974
22.2636 | -103.8115
22.2673 | -103.8241
22.271 | -103.8374
22.2747 | -103.8518
22.2783 | -103.8657
22.282 | -103.8779
22.2857 | -103.8908
22.2894 | -103.8999
22.293 | -103.9088
22.2967 | -103.9156
22.3004 | -103.9232
22.3041 | -103.9291
22.3077 | -103.9338
22.3114 | -103.9407
22.3151 | -103.9469
22.3187 | -103.951
22.3224 | -103.9543
22.3261 | -103.9596
22.3298 | -103.9649
22.3334 | -103.9681
22.3371 | -103.9716
22.3408 | -103.9762
22.3445 | -103.9778
22.3481 | -103.9803
22.3518 | -103.9833
22.3555 | -103.9848
22.3592 | -103.983
22.3628 | -103.9816
22.3665 | -103.9806
22.3702 | -103.9804
22.3738 | -103.9806
22.3775 | -103.9804
22.3812 | -103.9798
22.3849 | -103.9799
22.3885 | -103.9822
22.3922 | -103.9841
22.3959 | -103.9859
22.3996 | -103.9863
22.4032 | -103.9871
22.4069 | -103.9878
22.4106 | -103.9919
22.4143 | -103.9954
22.4179 | -104.0012
22.4216 | -104.005
22.4253 | -104.0092
22.4289 | -104.0149
22.4326 | -104.0203
22.4363 | -104.0257
22.44 | -104.0313
22.4436 | -104.0361
22.4473 | -104.0425
22.451 | -104.0489
22.4547 | -104.057
22.4583 | -104.0654
22.462 | -104.0734
22.4657 | -104.0826
22.4694 | -104.0928
22.473 | -104.1019
22.4767 | -104.1115
22.4804 | -104.1208
22.484 | -104.1326
22.4877 | -104.1449
22.4914 | -104.1567
22.4951 | -104.1682
22.4987 | -104.177
22.5024 | -104.1858
22.5061 | -104.1947
22.5098 | -104.2037
22.5134 | -104.2097
22.5171 | -104.2147
22.5208 | -104.2211
22.5245 | -104.2255
22.5281 | -104.2271
22.5318 | -104.2245
22.5355 | -104.2195
22.5391 | -104.217
22.5428 | -104.2154
22.5465 | -104.2069
22.5502 | -104.1953
22.5538 | -104.1801
22.5575 | -104.1618
22.5612 | -104.1394
22.5649 | -104.1132
22.5685 | -104.0853
22.5722 | -104.0555
22.5759 | -104.0225
22.5795 | -103.9892
22.5832 | -103.9553
22.5869 | -103.9184
22.5906 | -103.8817
22.5942 | -103.8438
22.5979 | -103.801
22.6016 | -103.7564
22.6053 | -103.711
22.6089 | -103.6656
22.6126 | -103.6189
22.6163 | -103.5715
22.62 | -103.5224
22.6236 | -103.4742
22.6273 | -103.4252
22.631 | -103.3735
22.6346 | -103.3219
22.6383 | -103.2692
22.642 | -103.2171
22.6457 | -103.1645
22.6493 | -103.1126
22.653 | -103.0615
22.6567 | -103.0118
22.6604 | -102.9642
22.664 | -102.9188
22.6677 | -102.875
22.6714 | -102.8317
22.6751 | -102.7887
22.6787 | -102.7501
22.6824 | -102.7143
22.6861 | -102.6805
22.6897 | -102.6484
22.6934 | -102.6181
22.6971 | -102.5899
22.7008 | -102.5649
22.7044 | -102.5426
22.7081 | -102.5222
22.7118 | -102.5045
22.7155 | -102.4893
22.7191 | -102.4764
22.7228 | -102.4652
22.7265 | -102.4574
22.7302 | -102.4512
22.7338 | -102.4472
22.7375 | -102.4444
22.7412 | -102.4449
22.7448 | -102.4477
22.7485 | -102.4506
22.7522 | -102.4572
22.7559 | -102.4674
22.7595 | -102.4801
22.7632 | -102.4939
22.7669 | -102.5065
22.7706 | -102.5201
22.7742 | -102.5342
22.7779 | -102.5504
22.7816 | -102.5688
22.7853 | -102.588
22.7889 | -102.6034
22.7926 | -102.6203
22.7963 | -102.6362
22.7999 | -102.6538
22.8036 | -102.671
22.8073 | -102.6882
22.811 | -102.7042
22.8146 | -102.7197
22.8183 | -102.7334
22.822 | -102.7464
22.8257 | -102.7603
22.8293 | -102.7731
22.833 | -102.7837
22.8367 | -102.7956
22.8404 | -102.8078
22.844 | -102.8185
22.8477 | -102.8285
22.8514 | -102.8394
22.855 | -102.8486
22.8587 | -102.8578
22.8624 | -102.8647
22.8661 | -102.8716
22.8697 | -102.8796
22.8734 | -102.8861
22.8771 | -102.8923
22.8808 | -102.9
22.8844 | -102.9093
22.8881 | -102.917
22.8918 | -102.9235
22.8955 | -102.9292
22.8991 | -102.9328
22.9028 | -102.9323
22.9065 | -102.9323
22.9101 | -102.9319
22.9138 | -102.9309
22.9175 | -102.9307
22.9212 | -102.9342
22.9248 | -102.938
22.9285 | -102.9395
22.9322 | -102.9405
22.9359 | -102.9415
22.9395 | -102.9437
22.9432 | -102.9449
22.9469 | -102.9426
22.9505 | -102.9437
22.9542 | -102.9468
22.9579 | -102.949
22.9616 | -102.9523
22.9652 | -102.9549
22.9689 | -102.9571
22.9726 | -102.9583
22.9763 | -102.9575
22.9799 | -102.957
22.9836 | -102.9567
22.9873 | -102.9569
22.991 | -102.9552
22.9946 | -102.9524
22.9983 | -102.9506
23.002 | -102.9446
23.0056 | -102.9392
23.0093 | -102.9357
23.013 | -102.9318
23.0167 | -102.9278
23.0203 | -102.9232
23.024 | -102.9169
23.0277 | -102.9106
23.0314 | -102.9055
23.035 | -102.9002
23.0387 | -102.897
23.0424 | -102.8931
23.0461 | -102.8898
23.0497 | -102.8851
23.0534 | -102.88
23.0571 | -102.8757
23.0607 | -102.8738
23.0644 | -102.8697
23.0681 | -102.8641
23.0718 | -102.862
23.0754 | -102.8585
23.0791 | -102.8531
23.0828 | -102.8476
23.0865 | -102.8435
23.0901 | -102.8424
23.0938 | -102.841
23.0975 | -102.8402
23.1012 | -102.8378
23.1048 | -102.8339
23.1085 | -102.8319
23.1122 | -102.8316
23.1158 | -102.8326
23.1195 | -102.8345
23.1232 | -102.8383
23.1269 | -102.8375
23.1305 | -102.8352
23.1342 | -102.8321
23.1379 | -102.8313
23.1416 | -102.8308
23.1452 | -102.8298
23.1489 | -102.8285
23.1526 | -102.8264
23.1563 | -102.8236
23.1599 | -102.8227
23.1636 | -102.8216
23.1673 | -102.8198
23.1709 | -102.8163
23.1746 | -102.8112
23.1783 | -102.8063
23.182 | -102.804
23.1856 | -102.8003
23.1893 | -102.7957
23.193 | -102.7893
23.1967 | -102.7821
23.2003 | -102.778
23.204 | -102.7737
23.2077 | -102.7679
23.2113 | -102.7594
23.215 | -102.7518
23.2187 | -102.7465
23.2224 | -102.7408
23.226 | -102.7334
23.2297 | -102.728
23.2334 | -102.7228
23.2371 | -102.718
23.2407 | -102.7113
23.2444 | -102.7039
23.2481 | -102.694
23.2518 | -102.6851
23.2554 | -102.6792
23.2591 | -102.6727
23.2628 | -102.6658
23.2664 | -102.6584
23.2701 | -102.6523
23.2738 | -102.6493
23.2775 | -102.6435
23.2811 | -102.6359
23.2848 | -102.6313
23.2885 | -102.6266
23.2922 | -102.6218
23.2958 | -102.6188
23.2995 | -102.6173
23.3032 | -102.6182
23.3069 | -102.6185
23.3105 | -102.619
23.3142 | -102.6294
23.3179 | -102.6292
23.3215 | -102.6314
23.3252 | -102.6362
23.3289 | -102.6409
23.3326 | -102.6483
23.3362 | -102.6542
23.3399 | -102.6599
23.3436 | -102.6675
23.3473 | -102.676
23.3509 | -102.6851
23.3546 | -102.6958
23.3583 | -102.7073
23.362 | -102.7174
23.3656 | -102.7282
23.3693 | -102.7379
23.373 | -102.7497
23.3766 | -102.7622
23.3803 | -102.7745
23.384 | -102.7868
23.3877 | -102.7992
23.3913 | -102.8107
23.395 | -102.8235
23.3987 | -102.8378
23.4024 | -102.8496
23.406 | -102.8609
23.4097 | -102.8727
23.4134 | -102.8855
23.4171 | -102.8975
23.4207 | -102.9078
23.4244 | -102.9174
23.4281 | -102.9261
23.4317 | -102.9355
23.4354 | -102.9449
23.4391 | -102.9521
23.4428 | -102.9586
23.4464 | -102.9659
23.4501 | -102.972
23.4538 | -102.9775
23.4575 | -102.9858
23.4611 | -102.9946
23.4648 | -103.001
23.4685 | -103.0054
23.4722 | -103.0101
23.4758 | -103.0161
23.4795 | -103.022
23.4832 | -103.0275
23.4868 | -103.0313
23.4905 | -103.0338
23.4942 | -103.0367
23.4979 | -103.0389
23.5015 | -103.0426
23.5052 | -103.048
23.5089 | -103.0529
23.5126 | -103.0571
23.5162 | -103.0625
23.5199 | -103.0663
23.5236 | -103.0702
23.5273 | -103.0725
23.5309 | -103.0714
23.5346 | -103.0723
23.5383 | -103.0743
23.5419 | -103.0746
23.5456 | -103.0747
23.5493 | -103.0751
23.553 | -103.0758
23.5566 | -103.0787
23.5603 | -103.0819
23.564 | -103.0836
23.5677 | -103.0845
23.5713 | -103.086
23.575 | -103.0878
23.5787 | -103.0901
23.5824 | -103.0909
23.586 | -103.0914
23.5897 | -103.0926
23.5934 | -103.096
23.597 | -103.0978
23.6007 | -103.0975
23.6044 | -103.0977
23.6081 | -103.1002
23.6117 | -103.1049
23.6154 | -103.1079
23.6191 | -103.1109
23.6228 | -103.1136
23.6264 | -103.1165
23.6301 | -103.1197
23.6338 | -103.1235
23.6374 | -103.1275
23.6411 | -103.1305
23.6448 | -103.1331
23.6485 | -103.1359
23.6521 | -103.1429
23.6558 | -103.1481
23.6595 | -103.152
23.6632 | -103.158
23.6668 | -103.1613
23.6705 | -103.1634
23.6742 | -103.1672
23.6779 | -103.1705
23.6815 | -103.1735
23.6852 | -103.1782
23.6889 | -103.18
23.6925 | -103.1798
23.6962 | -103.1823
23.6999 | -103.1865
23.7036 | -103.19
23.7072 | -103.1918
23.7109 | -103.1925
23.7146 | -103.1931
23.7183 | -103.1945
23.7219 | -103.1948
23.7256 | -103.195
23.7293 | -103.1964
23.733 | -103.1971
23.7366 | -103.1993
23.7403 | -103.2018
23.744 | -103.2025
23.7476 | -103.2013
23.7513 | -103.2008
23.755 | -103.1983
23.7587 | -103.1932
23.7623 | -103.1897
23.766 | -103.1864
23.7697 | -103.1839
23.7734 | -103.1818
23.777 | -103.18
23.7807 | -103.1753
23.7844 | -103.1714
23.7881 | -103.1653
23.7917 | -103.1576
23.7954 | -103.1521
23.7991 | -103.1496
23.8027 | -103.1467
23.8064 | -103.1434
23.8101 | -103.138
23.8138 | -103.1305
23.8174 | -103.1244
23.8211 | -103.1198
23.8248 | -103.1147
23.8285 | -103.1088
23.8321 | -103.1044
23.8358 | -103.1004
23.8395 | -103.0955
23.8432 | -103.0895
23.8468 | -103.082
23.8505 | -103.0739
23.8542 | -103.066
23.8578 | -103.0603
23.8615 | -103.0544
23.8652 | -103.0502
23.8689 | -103.047
23.8725 | -103.0413
23.8762 | -103.0338
23.8799 | -103.0282
23.8836 | -103.0231
23.8872 | -103.0169
23.8909 | -103.0107
23.8946 | -103.0036
23.8983 | -102.9957
23.9019 | -102.9889
23.9056 | -102.9841
23.9093 | -102.98
23.9129 | -102.9753
23.9166 | -102.9711
23.9203 | -102.9672
23.924 | -102.9601
23.9276 | -102.9551
23.9313 | -102.9507
23.935 | -102.9466
23.9387 | -102.9438
23.9423 | -102.9406
23.946 | -102.9343
23.9497 | -102.929
23.9534 | -102.9245
23.957 | -102.9213
23.9607 | -102.9187
23.9644 | -102.9152
23.968 | -102.9122
23.9717 | -102.9089
23.9754 | -102.9048
23.9791 | -102.9013
23.9827 | -102.8973
23.9864 | -102.8944
23.9901 | -102.8883
23.9938 | -102.8827
23.9974 | -102.8785
24.0011 | -102.8738
24.0048 | -102.8707
24.0084 | -102.8679
24.0121 | -102.8636
24.0158 | -102.8594
24.0195 | -102.8566
24.0231 | -102.8557
24.0268 | -102.8569
24.0305 | -102.8596
24.0342 | -102.861
24.0378 | -102.8614
24.0415 | -102.8628
24.0452 | -102.8652
24.0489 | -102.8621
24.0525 | -102.862
24.0562 | -102.8643
24.0599 | -102.867
24.0635 | -102.869
24.0672 | -102.8714
24.0709 | -102.8743
24.0746 | -102.8753
24.0782 | -102.8778
24.0819 | -102.8809
24.0856 | -102.8845
24.0893 | -102.8891
24.0929 | -102.896
24.0966 | -102.9042
24.1003 | -102.911
24.104 | -102.9147
24.1076 | -102.9163
24.1113 | -102.9154
24.115 | -102.9145
24.1186 | -102.9138
24.1223 | -102.9138
24.126 | -102.9118
24.1297 | -102.9098
24.1333 | -102.9065
24.137 | -102.902
24.1407 | -102.8985
24.1444 | -102.8948
24.148 | -102.891
24.1517 | -102.8882
24.1554 | -102.8829
24.1591 | -102.8763
24.1627 | -102.8671
24.1664 | -102.8583
24.1701 | -102.8488
24.1737 | -102.8373
24.1774 | -102.8264
24.1811 | -102.8157
24.1848 | -102.8058
24.1884 | -102.7956
24.1921 | -102.7809
24.1958 | -102.7643
24.1995 | -102.7495
24.2031 | -102.7357
24.2068 | -102.7211
24.2105 | -102.7055
24.2142 | -102.6904
24.2178 | -102.6758
24.2215 | -102.6605
24.2252 | -102.6444
24.2288 | -102.6268
24.2325 | -102.6083
24.2362 | -102.5901
24.2399 | -102.5716
24.2435 | -102.5534
24.2472 | -102.5358
24.2509 | -102.5185
24.2546 | -102.5013
24.2582 | -102.4861
24.2619 | -102.4694
24.2656 | -102.455
24.2692 | -102.4396
24.2729 | -102.4221
24.2766 | -102.4053
24.2803 | -102.3858
24.2839 | -102.3604
24.2876 | -102.3363
24.2913 | -102.3109
24.295 | -102.2904
24.2986 | -102.2736
24.3023 | -102.2577
24.306 | -102.2415
24.3097 | -102.225
24.3133 | -102.209
24.317 | -102.1937
24.3207 | -102.1803
24.3243 | -102.1683
24.328 | -102.1545
24.3317 | -102.143
24.3354 | -102.1308
24.339 | -102.1182
24.3427 | -102.1059
24.3464 | -102.0935
24.3501 | -102.0811
24.3537 | -102.0684
24.3574 | -102.056
24.3611 | -102.0433
24.3648 | -102.0284
24.3684 | -102.0154
24.3721 | -102.0029
24.3758 | -101.996
24.3794 | -101.9854
24.3831 | -101.9743
24.3868 | -101.9629
24.3905 | -101.9532
24.3941 | -101.9489
24.3978 | -101.9439
24.4015 | -101.9374
24.4052 | -101.9313
24.4088 | -101.9271
24.4125 | -101.924
24.4162 | -101.9217
24.4199 | -101.9219
24.4235 | -101.9234
24.4272 | -101.9263
24.4309 | -101.9297
24.4345 | -101.934
24.4382 | -101.94
24.4419 | -101.9486
24.4456 | -101.9587
24.4492 | -101.9717
24.4529 | -101.9863
24.4566 | -102.0053
24.4603 | -102.0256
24.4639 | -102.0467
24.4676 | -102.0678
24.4713 | -102.0928
24.475 | -102.1195
24.4786 | -102.1462
24.4823 | -102.1733
24.486 | -102.2017
24.4896 | -102.2317
24.4933 | -102.2639
24.497 | -102.2975
24.5007 | -102.333
24.5043 | -102.3708
24.508 | -102.4112
24.5117 | -102.4518
24.5154 | -102.4918
24.519 | -102.5309
24.5227 | -102.5694
24.5264 | -102.6099
24.5301 | -102.6499
24.5337 | -102.6915
24.5374 | -102.7366
24.5411 | -102.7806
24.5447 | -102.8215
24.5484 | -102.8601
24.5521 | -102.8972
24.5558 | -102.9337
24.5594 | -102.9668
24.5631 | -102.9931
24.5668 | -103.0135
24.5705 | -103.0321
24.5741 | -103.0579
24.5778 | -103.0807
24.5815 | -103.0953
24.5852 | -103.0999
24.5888 | -103.0622
24.5925 | -102.974
24.5962 | -102.8767
24.5998 | -102.8237
24.6035 | -102.783
24.6072 | -102.7411
24.6109 | -102.6925
24.6145 | -102.6344
24.6182 | -102.5683
24.6219 | -102.4996
24.6256 | -102.4303
24.6292 | -102.3614
24.6329 | -102.2862
24.6366 | -102.2053
24.6403 | -102.1137
24.6439 | -102.0152
24.6476 | -101.914
24.6513 | -101.8101
24.6549 | -101.7028
24.6586 | -101.5926
24.6623 | -101.4813
24.666 | -101.366
24.6696 | -101.2511
24.6733 | -101.1377
24.677 | -101.0242
24.6807 | -100.9127
24.6843 | -100.8026
24.688 | -100.6929
24.6917 | -100.5844
24.6953 | -100.478
24.699 | -100.3746
24.7027 | -100.275
24.7064 | -100.1794
24.71 | -100.0886
24.7137 | -100.0018
24.7174 | -99.9202
24.7211 | -99.8429
24.7247 | -99.7702
24.7284 | -99.7039
24.7321 | -99.6436
24.7358 | -99.5891
24.7394 | -99.5401
24.7431 | -99.4962
24.7468 | -99.458
24.7504 | -99.4264
24.7541 | -99.4019
24.7578 | -99.3831
24.7615 | -99.3703
24.7651 | -99.365
24.7688 | -99.3662
24.7725 | -99.3733
24.7762 | -99.3876
24.7798 | -99.4085
24.7835 | -99.4358
24.7872 | -99.4694
24.7909 | -99.5086
24.7945 | -99.5538
24.7982 | -99.6053
24.8019 | -99.6622
24.8055 | -99.7239
24.8092 | -99.7918
24.8129 | -99.8657
24.8166 | -99.9443
24.8202 | -100.0277
24.8239 | -100.1156
24.8276 | -100.2069
24.8313 | -100.3033
24.8349 | -100.4028
24.8386 | -100.5053
24.8423 | -100.6105
24.846 | -100.7173
24.8496 | -100.8268
24.8533 | -100.9381
24.857 | -101.0503
24.8606 | -101.1627
24.8643 | -101.2767
24.868 | -101.3933
24.8717 | -101.5092
24.8753 | -101.6234
24.879 | -101.7369
24.8827 | -101.8477
24.8864 | -101.9563
24.89 | -102.0618
24.8937 | -102.1645
24.8974 | -102.2663
24.9011 | -102.3664
24.9047 | -102.4144
24.9084 | -102.505
24.9121 | -102.5942
24.9157 | -102.6849
24.9194 | -102.7741
24.9231 | -102.855
24.9268 | -102.9241
24.9304 | -102.9928
24.9341 | -103.0611
24.9378 | -103.1317
24.9415 | -103.2299
24.9451 | -103.358
24.9488 | -103.4752
24.9525 | -103.5245
24.9562 | -103.5441
24.9598 | -103.5599
24.9635 | -103.5754
24.9672 | -103.5932
24.9708 | -103.6143
24.9745 | -103.6288
24.9782 | -103.6376
24.9819 | -103.6421
24.9855 | -103.6447
24.9892 | -103.6467
24.9929 | -103.6487
24.9966 | -103.6506
25.0002 | -103.6515
25.0039 | -103.6512
25.0076 | -103.6534
25.0112 | -103.6564
25.0149 | -103.6577
25.0186 | -103.6623
25.0223 | -103.6641
25.0259 | -103.6617
25.0296 | -103.6599
25.0333 | -103.6573
25.037 | -103.6531
25.0406 | -103.6521
25.0443 | -103.6529
25.048 | -103.6523
25.0517 | -103.6523
25.0553 | -103.6529
25.059 | -103.65
25.0627 | -103.6479
25.0663 | -103.6454
25.07 | -103.6423
25.0737 | -103.6433
25.0774 | -103.6459
25.081 | -103.6464
25.0847 | -103.6475
25.0884 | -103.6504
25.0921 | -103.6533
25.0957 | -103.658
25.0994 | -103.6639
25.1031 | -103.6688
25.1068 | -103.6729
25.1104 | -103.6779
25.1141 | -103.684
25.1178 | -103.6896
25.1214 | -103.6949
25.1251 | -103.6986
25.1288 | -103.7023
25.1325 | -103.7077
25.1361 | -103.7121
25.1398 | -103.7154
25.1435 | -103.721
25.1472 | -103.7259
25.1508 | -103.7288
25.1545 | -103.7349
25.1582 | -103.7412
25.1619 | -103.7459
25.1655 | -103.7498
25.1692 | -103.7539
25.1729 | -103.7584
25.1765 | -103.7614
25.1802 | -103.7669
25.1839 | -103.7711
25.1876 | -103.7744
25.1912 | -103.7772
25.1949 | -103.7787
25.1986 | -103.7816
25.2023 | -103.7836
25.2059 | -103.7849
25.2096 | -103.7853
25.2133 | -103.7851
25.217 | -103.784
25.2206 | -103.7839
25.2243 | -103.784
25.228 | -103.7808
25.2316 | -103.7757
25.2353 | -103.7725
25.239 | -103.7672
25.2427 | -103.7607
25.2463 | -103.7541
25.25 | -103.7475
25.2537 | -103.7409
25.2574 | -103.734
25.261 | -103.7278
25.2647 | -103.7223
25.2684 | -103.7147
25.2721 | -103.7047
25.2757 | -103.6947
25.2794 | -103.6861
25.2831 | -103.6747
25.2867 | -103.6611
25.2904 | -103.6476
25.2941 | -103.6339
25.2978 | -103.6194
25.3014 | -103.6034
25.3051 | -103.5867
25.3088 | -103.5698
25.3125 | -103.5541
25.3161 | -103.5376
25.3198 | -103.5182
25.3235 | -103.4969
25.3271 | -103.4744
25.3308 | -103.4529
25.3345 | -103.4327
25.3382 | -103.4138
25.3418 | -103.394
25.3455 | -103.3745
25.3492 | -103.3571
25.3529 | -103.3378
25.3565 | -103.3184
25.3602 | -103.301
25.3639 | -103.2808
25.3676 | -103.2612
25.3712 | -103.2431
25.3749 | -103.2248
25.3786 | -103.2075
25.3822 | -103.191
25.3859 | -103.1744
25.3896 | -103.1585
25.3933 | -103.1417
25.3969 | -103.1249
25.4006 | -103.1094
25.4043 | -103.0946
25.408 | -103.0781
25.4116 | -103.0585
25.4153 | -103.0427
25.419 | -103.028
25.4227 | -103.0136
25.4263 | -102.9991
25.43 | -102.9837
25.4337 | -102.9725
25.4373 | -102.9581
25.441 | -102.9453
25.4447 | -102.9337
25.4484 | -102.923
25.452 | -102.9116
25.4557 | -102.901
25.4594 | -102.8915
25.4631 | -102.8837
25.4667 | -102.8762
25.4704 | -102.8694
25.4741 | -102.8626
25.4778 | -102.8571
25.4814 | -102.8518
25.4851 | -102.8471
25.4888 | -102.8441
25.4924 | -102.841
25.4961 | -102.8375
25.4998 | -102.8344
25.5035 | -102.8317
25.5071 | -102.8299
25.5108 | -102.8287
25.5145 | -102.8285
25.5182 | -102.829
25.5218 | -102.8291
25.5255 | -102.8281
25.5292 | -102.8295
25.5329 | -102.8315
25.5365 | -102.8347
25.5402 | -102.8379
25.5439 | -102.8405
25.5475 | -102.8441
25.5512 | -102.8473
25.5549 | -102.8501
25.5586 | -102.8542
25.5622 | -102.8587
25.5659 | -102.8621
25.5696 | -102.8672
25.5733 | -102.8725
25.5769 | -102.8765
25.5806 | -102.8803
25.5843 | -102.8842
25.588 | -102.8861
25.5916 | -102.8891
25.5953 | -102.8912
25.599 | -102.8946
25.6026 | -102.9005
25.6063 | -102.9064
25.61 | -102.9097
25.6137 | -102.9129
25.6173 | -102.9159
25.621 | -102.918
25.6247 | -102.9199
25.6284 | -102.9215
25.632 | -102.9226
25.6357 | -102.924
25.6394 | -102.9285
25.6431 | -102.9337
25.6467 | -102.9383
25.6504 | -102.9432
25.6541 | -102.9468
25.6577 | -102.9507
25.6614 | -102.9579
25.6651 | -102.9622
25.6688 | -102.9654
25.6724 | -102.9709
25.6761 | -102.9769
25.6798 | -102.9826
25.6835 | -102.9907
25.6871 | -102.9988
25.6908 | -103.0071
25.6945 | -103.014
25.6981 | -103.0203
25.7018 | -103.0267
25.7055 | -103.034
25.7092 | -103.04
25.7128 | -103.0478
25.7165 | -103.0566
25.7202 | -103.0645
25.7239 | -103.0746
25.7275 | -103.0849
25.7312 | -103.0943
25.7349 | -103.1023
25.7386 | -103.1118
25.7422 | -103.1219
25.7459 | -103.1297
25.7496 | -103.1367
25.7532 | -103.1472
25.7569 | -103.1557
25.7606 | -103.164
25.7643 | -103.1745
25.7679 | -103.1845
25.7716 | -103.1964
25.7753 | -103.2077
25.779 | -103.2163
25.7826 | -103.2213
25.7863 | -103.2285
25.79 | -103.2361
25.7937 | -103.2446
25.7973 | -103.252
25.801 | -103.2578
25.8047 | -103.2628
25.8083 | -103.2682
25.812 | -103.2738
25.8157 | -103.2811
25.8194 | -103.288
25.823 | -103.2917
25.8267 | -103.2915
25.8304 | -103.2943
25.8341 | -103.2966
25.8377 | -103.2973
25.8414 | -103.2979
25.8451 | -103.2998
25.8488 | -103.3033
25.8524 | -103.3048
25.8561 | -103.306
25.8598 | -103.306
25.8634 | -103.3058
25.8671 | -103.3054
25.8708 | -103.3006
25.8745 | -103.2971
25.8781 | -103.2927
25.8818 | -103.2867
25.8855 | -103.2814
25.8892 | -103.2781
25.8928 | -103.2736
25.8965 | -103.2676
25.9002 | -103.2603
25.9039 | -103.2508
25.9075 | -103.2425
25.9112 | -103.2344
25.9149 | -103.2278
25.9185 | -103.2216
25.9222 | -103.2132
25.9259 | -103.204
25.9296 | -103.1965
25.9332 | -103.1894
25.9369 | -103.1816
25.9406 | -103.1725
25.9443 | -103.1638
25.9479 | -103.1552
25.9516 | -103.1496
25.9553 | -103.1441
25.959 | -103.1369
25.9626 | -103.1286
25.9663 | -103.1232
25.97 | -103.114
25.9736 | -103.1061
25.9773 | -103.0977
25.981 | -103.0877
25.9847 | -103.0821
25.9883 | -103.0764
25.992 | -103.0695
25.9957 | -103.0658
25.9994 | -103.0646
26.003 | -103.0628
26.0067 | -103.0599
26.0104 | -103.0568
26.0141 | -103.0513
26.0177 | -103.0482
26.0214 | -103.0473
26.0251 | -103.0478
26.0287 | -103.047
26.0324 | -103.0479
26.0361 | -103.0506
26.0398 | -103.052
26.0434 | -103.0525
26.0471 | -103.0541
26.0508 | -103.0557
26.0545 | -103.0579
26.0581 | -103.0597
26.0618 | -103.0624
26.0655 | -103.0661
26.0691 | -103.069
26.0728 | -103.0698
26.0765 | -103.0735
26.0802 | -103.0769
26.0838 | -103.0782
26.0875 | -103.0813
26.0912 | -103.086
26.0949 | -103.0878
26.0985 | -103.0884
26.1022 | -103.0891
26.1059 | -103.09
26.1096 | -103.0931
26.1132 | -103.0986
26.1169 | -103.1032
26.1206 | -103.1053
26.1242 | -103.1071
26.1279 | -103.1087
26.1316 | -103.1107
26.1353 | -103.1144
26.1389 | -103.1166
26.1426 | -103.1176
26.1463 | -103.1198
26.15 | -103.1253
26.1536 | -103.1284
26.1573 | -103.13
26.161 | -103.1322
26.1647 | -103.134
26.1683 | -103.1362
26.172 | -103.1368
26.1757 | -103.1382
26.1793 | -103.14
26.183 | -103.1422
26.1867 | -103.144
26.1904 | -103.1474
26.194 | -103.1539
26.1977 | -103.1567
26.2014 | -103.1605
26.2051 | -103.1648
26.2087 | -103.1675
26.2124 | -103.171
26.2161 | -103.1739
26.2198 | -103.1767
26.2234 | -103.1804
26.2271 | -103.185
26.2308 | -103.1892
26.2344 | -103.1928
26.2381 | -103.1996
26.2418 | -103.2064
26.2455 | -103.2129
26.2491 | -103.2203
26.2528 | -103.2283
26.2565 | -103.2355
26.2602 | -103.2421
26.2638 | -103.2511
26.2675 | -103.2627
26.2712 | -103.2741
26.2749 | -103.2843
26.2785 | -103.2946
26.2822 | -103.3058
26.2859 | -103.3171
26.2895 | -103.327
26.2932 | -103.3384
26.2969 | -103.3488
26.3006 | -103.3598
26.3042 | -103.3731
26.3079 | -103.3878
26.3116 | -103.4014
26.3153 | -103.415
26.3189 | -103.4286
26.3226 | -103.4403
26.3263 | -103.4538
26.33 | -103.4682
26.3336 | -103.482
26.3373 | -103.496
26.341 | -103.5095
26.3446 | -103.5204
26.3483 | -103.533
26.352 | -103.5471
26.3557 | -103.556
26.3593 | -103.5673
26.363 | -103.5799
26.3667 | -103.5879
26.3704 | -103.593
26.374 | -103.5967
26.3777 | -103.6003
26.3814 | -103.6038
26.385 | -103.611
26.3887 | -103.6165
26.3924 | -103.6199
26.3961 | -103.6206
26.3997 | -103.6191
26.4034 | -103.6172
26.4071 | -103.6178
26.4108 | -103.6185
26.4144 | -103.618
26.4181 | -103.617
26.4218 | -103.6145
26.4255 | -103.6113
26.4291 | -103.6084
26.4328 | -103.6038
26.4365 | -103.5973
26.4401 | -103.5935
26.4438 | -103.5897
26.4475 | -103.5827
26.4512 | -103.5751
26.4548 | -103.5677
26.4585 | -103.5567
26.4622 | -103.5454
26.4659 | -103.5361
26.4695 | -103.5288
26.4732 | -103.522
26.4769 | -103.513
26.4806 | -103.5008
26.4842 | -103.4889
26.4879 | -103.477
26.4916 | -103.4641
26.4952 | -103.454
26.4989 | -103.4441
26.5026 | -103.433
26.5063 | -103.4196
26.5099 | -103.4062
26.5136 | -103.3962
26.5173 | -103.3875
26.521 | -103.3783
26.5246 | -103.3694
26.5283 | -103.361
26.532 | -103.3525
26.5357 | -103.3454
26.5393 | -103.3397
26.543 | -103.3332
26.5467 | -103.33
26.5503 | -103.3262
26.554 | -103.3211
26.5577 | -103.3169
26.5614 | -103.3147
26.565 | -103.3114
26.5687 | -103.3081
26.5724 | -103.3072
26.5761 | -103.3062
26.5797 | -103.3054
26.5834 | -103.306
26.5871 | -103.306
26.5908 | -103.3074
26.5944 | -103.3116
26.5981 | -103.3145
26.6018 | -103.3189
26.6054 | -103.3248
26.6091 | -103.3297
26.6128 | -103.3338
26.6165 | -103.3389
26.6201 | -103.3448
26.6238 | -103.3504
26.6275 | -103.356
26.6312 | -103.3624
26.6348 | -103.3691
26.6385 | -103.3756
26.6422 | -103.3832
26.6459 | -103.3923
26.6495 | -103.4023
26.6532 | -103.4123
26.6569 | -103.4194
26.6605 | -103.4281
26.6642 | -103.4396
26.6679 | -103.4485
26.6716 | -103.4582
26.6752 | -103.4675
26.6789 | -103.4752
26.6826 | -103.4835
26.6863 | -103.4933
26.6899 | -103.5007
26.6936 | -103.5074
26.6973 | -103.518
26.7009 | -103.5268
26.7046 | -103.5368
26.7083 | -103.5452
26.712 | -103.5545
26.7156 | -103.565
26.7193 | -103.5737
26.723 | -103.582
26.7267 | -103.5894
26.7303 | -103.5983
26.734 | -103.604
26.7377 | -103.6078
26.7414 | -103.6142
26.745 | -103.6224
26.7487 | -103.6302
26.7524 | -103.6377
26.756 | -103.6439
26.7597 | -103.6468
26.7634 | -103.6497
26.7671 | -103.6536
26.7707 | -103.6581
26.7744 | -103.665
26.7781 | -103.6696
26.7818 | -103.6687
26.7854 | -103.6692
26.7891 | -103.6706
26.7928 | -103.67
26.7965 | -103.6697
26.8001 | -103.6685
26.8038 | -103.6652
26.8075 | -103.6623
26.8111 | -103.6615
26.8148 | -103.6591
26.8185 | -103.6563
26.8222 | -103.6534
26.8258 | -103.6509
26.8295 | -103.6493
26.8332 | -103.6482
26.8369 | -103.6436
26.8405 | -103.6385
26.8442 | -103.6341
26.8479 | -103.6296
26.8516 | -103.6247
26.8552 | -103.6195
26.8589 | -103.6126
26.8626 | -103.6053
26.8662 | -103.601
26.8699 | -103.5951
26.8736 | -103.5901
26.8773 | -103.5859
26.8809 | -103.582
26.8846 | -103.5789
26.8883 | -103.5764
26.892 | -103.5723
26.8956 | -103.5661
26.8993 | -103.563
26.903 | -103.5595
26.9067 | -103.556
26.9103 | -103.5549
26.914 | -103.5544
26.9177 | -103.5531
26.9213 | -103.5547
26.925 | -103.5553
26.9287 | -103.5571
26.9324 | -103.5605
26.936 | -103.5618
26.9397 | -103.5598
26.9434 | -103.5629
26.9471 | -103.5682
26.9507 | -103.573
26.9544 | -103.5772
26.9581 | -103.5839
26.9618 | -103.5904
26.9654 | -103.5967
26.9691 | -103.6036
26.9728 | -103.6088
26.9764 | -103.6153
26.9801 | -103.6216
26.9838 | -103.6273
26.9875 | -103.636
26.9911 | -103.6465
26.9948 | -103.6541
26.9985 | -103.662
27.0022 | -103.6706
27.0058 | -103.6777
27.0095 | -103.6859
27.0132 | -103.6956
27.0169 | -103.7029
27.0205 | -103.7067
27.0242 | -103.7076
27.0279 | -103.7139
27.0315 | -103.7189
27.0352 | -103.7239
27.0389 | -103.7293
27.0426 | -103.7319
27.0462 | -103.7345
27.0499 | -103.7379
27.0536 | -103.7412
27.0573 | -103.7435
27.0609 | -103.7445
27.0646 | -103.7451
27.0683 | -103.7428
27.072 | -103.7423
27.0756 | -103.7436
27.0793 | -103.7396
27.083 | -103.7347
27.0866 | -103.7303
27.0903 | -103.7256
27.094 | -103.7203
27.0977 | -103.7153
27.1013 | -103.7114
27.105 | -103.7073
27.1087 | -103.7029
27.1124 | -103.6979
27.116 | -103.6942
27.1197 | -103.6899
27.1234 | -103.6847
27.127 | -103.6787
27.1307 | -103.6744
27.1344 | -103.6687
27.1381 | -103.6632
27.1417 | -103.6597
27.1454 | -103.6551
27.1491 | -103.648
27.1528 | -103.6419
27.1564 | -103.6361
27.1601 | -103.6326
27.1638 | -103.6316
27.1675 | -103.6298
27.1711 | -103.6277
27.1748 | -103.6268
27.1785 | -103.6277
27.1821 | -103.6289
27.1858 | -103.6321
27.1895 | -103.6349
27.1932 | -103.6355
27.1968 | -103.6378
27.2005 | -103.6414
27.2042 | -103.6451
27.2079 | -103.6499
27.2115 | -103.655
27.2152 | -103.6618
27.2189 | -103.6687
27.2226 | -103.6759
27.2262 | -103.6833
27.2299 | -103.6898
27.2336 | -103.6967
27.2372 | -103.7057
27.2409 | -103.7148
27.2446 | -103.7255
27.2483 | -103.7346
27.2519 | -103.7426
27.2556 | -103.7479
27.2593 | -103.7536
27.263 | -103.7609
27.2666 | -103.7693
27.2703 | -103.7762
27.274 | -103.7826
27.2777 | -103.7919
27.2813 | -103.7989
27.285 | -103.8052
27.2887 | -103.8114
27.2923 | -103.8182
27.296 | -103.8231
27.2997 | -103.8289
27.3034 | -103.8343
27.307 | -103.8402
27.3107 | -103.8444
27.3144 | -103.8485
27.3181 | -103.8556
27.3217 | -103.8641
27.3254 | -103.8691
27.3291 | -103.8713
27.3328 | -103.873
27.3364 | -103.8764
27.3401 | -103.8791
27.3438 | -103.8801
27.3474 | -103.8817
27.3511 | -103.8833
27.3548 | -103.885
27.3585 | -103.8874
27.3621 | -103.8899
27.3658 | -103.8918
27.3695 | -103.8935
27.3732 | -103.8944
27.3768 | -103.8972
27.3805 | -103.8984
27.3842 | -103.8992
27.3879 | -103.9006
27.3915 | -103.9016
27.3952 | -103.9015
27.3989 | -103.9049
27.4025 | -103.9082
27.4062 | -103.9104
27.4099 | -103.9094
27.4136 | -103.9101
27.4172 | -103.9128
27.4209 | -103.914
27.4246 | -103.9166
27.4283 | -103.921
27.4319 | -103.9235
27.4356 | -103.9258
27.4393 | -103.9286
27.4429 | -103.9305
27.4466 | -103.9336
27.4503 | -103.9369
27.454 | -103.9398
27.4576 | -103.943
27.4613 | -103.9452
27.465 | -103.9497
27.4687 | -103.9543
27.4723 | -103.9571
27.476 | -103.9602
27.4797 | -103.9615
27.4834 | -103.9642
27.487 | -103.9676
27.4907 | -103.9699
27.4944 | -103.9731
27.498 | -103.9752
27.5017 | -103.9766
27.5054 | -103.9804
27.5091 | -103.9824
27.5127 | -103.9841
27.5164 | -103.9857
27.5201 | -103.9879
27.5238 | -103.9911
27.5274 | -103.9944
27.5311 | -103.9984
27.5348 | -104.0034
27.5385 | -104.0078
27.5421 | -104.0102
27.5458 | -104.0104
27.5495 | -104.0103
27.5531 | -104.0082
27.5568 | -104.0091
27.5605 | -104.0096
27.5642 | -104.0112
27.5678 | -104.0103
27.5715 | -104.0098
27.5752 | -104.0077
27.5789 | -104.0054
27.5825 | -104.0064
27.5862 | -104.0061
27.5899 | -104.0048
27.5936 | -104.0049
27.5972 | -104.0045
27.6009 | -104.0016
27.6046 | -103.9981
27.6082 | -103.9961
27.6119 | -103.995
27.6156 | -103.9912
27.6193 | -103.9879
27.6229 | -103.9847
27.6266 | -103.9822
27.6303 | -103.9782
27.634 | -103.9708
27.6376 | -103.9626
27.6413 | -103.9561
27.645 | -103.9509
27.6487 | -103.9442
27.6523 | -103.9355
27.656 | -103.9265
27.6597 | -103.9184
27.6633 | -103.9107
27.667 | -103.901
27.6707 | -103.8917
27.6744 | -103.8849
27.678 | -103.8775
27.6817 | -103.8681
27.6854 | -103.8597
27.6891 | -103.8511
27.6927 | -103.8414
27.6964 | -103.831
27.7001 | -103.8216
27.7038 | -103.8124
27.7074 | -103.8021
27.7111 | -103.7929
27.7148 | -103.7828
27.7184 | -103.7744
27.7221 | -103.7662
27.7258 | -103.7562
27.7295 | -103.7461
27.7331 | -103.7361
27.7368 | -103.7265
27.7405 | -103.7162
27.7442 | -103.7063
27.7478 | -103.6967
27.7515 | -103.6872
27.7552 | -103.6787
27.7588 | -103.6703
27.7625 | -103.6632
27.7662 | -103.6556
27.7699 | -103.6486
27.7735 | -103.6425
27.7772 | -103.6382
27.7809 | -103.6337
27.7846 | -103.63
27.7882 | -103.6261
27.7919 | -103.6244
27.7956 | -103.623
27.7993 | -103.623
27.8029 | -103.6213
27.8066 | -103.6194
27.8103 | -103.6206
27.8139 | -103.6221
27.8176 | -103.6239
27.8213 | -103.6261
27.825 | -103.6299
27.8286 | -103.6331
27.8323 | -103.6378
27.836 | -103.6433
27.8397 | -103.6488
27.8433 | -103.6566
27.847 | -103.6654
27.8507 | -103.6749
27.8544 | -103.6845
27.858 | -103.6942
27.8617 | -103.705
27.8654 | -103.718
27.869 | -103.7325
27.8727 | -103.7455
27.8764 | -103.7572
27.8801 | -103.7699
27.8837 | -103.7828
27.8874 | -103.7977
27.8911 | -103.813
27.8948 | -103.829
27.8984 | -103.8461
27.9021 | -103.8617
27.9058 | -103.8764
27.9095 | -103.8908
27.9131 | -103.9049
27.9168 | -103.9178
27.9205 | -103.9299
27.9241 | -103.9403
27.9278 | -103.9512
27.9315 | -103.9631
27.9352 | -103.9757
27.9388 | -103.9889
27.9425 | -104.0023
27.9462 | -104.0122
27.9499 | -104.0227
27.9535 | -104.0326
27.9572 | -104.0419
27.9609 | -104.0515
27.9646 | -104.0594
27.9682 | -104.0685
27.9719 | -104.0779
27.9756 | -104.0827
27.9792 | -104.0879
27.9829 | -104.0924
27.9866 | -104.0955
27.9903 | -104.0992
27.9939 | -104.1028
27.9976 | -104.1054
28.0013 | -104.108