## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 233 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 235
Hood | 02-92B/P | 174
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1007.8 | hPa |  | PRC PRO 01
Humidity | 0 | %/Hr
Internal temperature of calibration bench | 24.08 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1027 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.94 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.03 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3061 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 28/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 29/05/2018
Engineer | Anthony HUE | Approval |  | 29/05/2018
CEO | Robert GLEMOT | Approval |  | 29/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3722 | 57.3563
0.012209 | -35.9464 | 38.139
0.018313 | -34.8857 | 28.0528
0.024417 | -34.4292 | 21.829
0.030521 | -34.3273 | 18.1053
0.036626 | -34.1862 | 15.0514
0.04273 | -34.0919 | 12.6768
0.048834 | -34.0467 | 11.1759
0.054938 | -33.9795 | 10.0547
0.061043 | -33.9161 | 8.9706
0.067147 | -33.9072 | 8.1504
0.073251 | -33.9103 | 7.3908
0.079355 | -33.9315 | 6.7282
0.08546 | -33.9473 | 6.1314
0.091564 | -33.9425 | 5.6958
0.097668 | -33.9234 | 5.4629
0.10377 | -33.9004 | 5.2206
0.10988 | -33.8826 | 4.9458
0.11598 | -33.8733 | 4.6527
0.12209 | -33.8758 | 4.3688
0.12819 | -33.8869 | 4.0845
0.13429 | -33.8989 | 3.7938
0.1404 | -33.9024 | 3.5106
0.1465 | -33.9044 | 3.2495
0.15261 | -33.9014 | 3.0273
0.15871 | -33.8943 | 2.8661
0.16482 | -33.8857 | 2.7671
0.17092 | -33.8836 | 2.734
0.17702 | -33.8816 | 2.6656
0.18313 | -33.881 | 2.5902
0.18923 | -33.8809 | 2.4814
0.19534 | -33.8798 | 2.3534
0.20144 | -33.8786 | 2.2083
0.20754 | -33.8755 | 2.0856
0.21365 | -33.8746 | 1.9696
0.21975 | -33.8741 | 1.8715
0.22586 | -33.8753 | 1.7846
0.23196 | -33.878 | 1.7137
0.23807 | -33.8823 | 1.6527
0.24417 | -33.8867 | 1.5947
0.25027 | -33.891 | 1.5329
0.25638 | -33.8936 | 1.471
0.26248 | -33.8956 | 1.4038
0.26859 | -33.8966 | 1.3377
0.27469 | -33.8972 | 1.2654
0.2808 | -33.8984 | 1.1908
0.2869 | -33.8994 | 1.1191
0.293 | -33.9006 | 1.0514
0.29911 | -33.9017 | 0.99143
0.30521 | -33.9029 | 0.93242
0.31132 | -33.9037 | 0.89525
0.31742 | -33.9051 | 0.8536
0.32353 | -33.9055 | 0.82016
0.32963 | -33.9065 | 0.79301
0.33573 | -33.9076 | 0.77487
0.34184 | -33.9091 | 0.75536
0.34794 | -33.9115 | 0.73267
0.35405 | -33.9138 | 0.70981
0.36015 | -33.9165 | 0.68547
0.36626 | -33.9187 | 0.65868
0.37236 | -33.921 | 0.62191
0.37846 | -33.9234 | 0.57697
0.38457 | -33.9244 | 0.5182
0.39067 | -33.9255 | 0.47068
0.39678 | -33.9262 | 0.41757
0.40288 | -33.9267 | 0.36211
0.40899 | -33.9265 | 0.30881
0.41509 | -33.9258 | 0.25664
0.42119 | -33.9257 | 0.21242
0.4273 | -33.9256 | 0.16507
0.4334 | -33.9255 | 0.11777
0.43951 | -33.9258 | 0.073241
0.44561 | -33.9267 | 0.032595
0.45172 | -33.9284 | 0.00027859
0.45782 | -33.9293 | -0.037621
0.46392 | -33.9306 | -0.074646
0.47003 | -33.9318 | -0.10763
0.47613 | -33.9328 | -0.14243
0.48224 | -33.9334 | -0.17806
0.48834 | -33.9324 | -0.21364
0.49445 | -33.9322 | -0.24977
0.50055 | -33.9317 | -0.2825
0.50665 | -33.9309 | -0.31758
0.51276 | -33.9299 | -0.35251
0.51886 | -33.9288 | -0.38724
0.52497 | -33.9275 | -0.41163
0.53107 | -33.9267 | -0.4417
0.53717 | -33.9259 | -0.47038
0.54328 | -33.925 | -0.49336
0.54938 | -33.9249 | -0.51596
0.55549 | -33.9249 | -0.53689
0.56159 | -33.9261 | -0.54789
0.5677 | -33.9264 | -0.56205
0.5738 | -33.9276 | -0.57806
0.5799 | -33.9284 | -0.5926
0.58601 | -33.9296 | -0.60198
0.59211 | -33.9314 | -0.60724
0.59822 | -33.9329 | -0.61589
0.60432 | -33.9345 | -0.62734
0.61043 | -33.9359 | -0.63432
0.61653 | -33.9372 | -0.6438
0.62263 | -33.9389 | -0.65231
0.62874 | -33.9396 | -0.66215
0.63484 | -33.9405 | -0.67057
0.64095 | -33.9411 | -0.67997
0.64705 | -33.9415 | -0.6889
0.65316 | -33.9417 | -0.70075
0.65926 | -33.9418 | -0.71627
0.66536 | -33.9406 | -0.73365
0.67147 | -33.9402 | -0.74645
0.67757 | -33.9399 | -0.76044
0.68368 | -33.9391 | -0.77939
0.68978 | -33.9383 | -0.79653
0.69589 | -33.9373 | -0.81562
0.70199 | -33.9363 | -0.83892
0.70809 | -33.9354 | -0.85829
0.7142 | -33.9347 | -0.87903
0.7203 | -33.9335 | -0.89847
0.72641 | -33.9328 | -0.92143
0.73251 | -33.9322 | -0.9418
0.73862 | -33.9324 | -0.96409
0.74472 | -33.9321 | -0.98445
0.75082 | -33.9318 | -1.0071
0.75693 | -33.9322 | -1.0294
0.76303 | -33.9328 | -1.0519
0.76914 | -33.9348 | -1.0727
0.77524 | -33.9358 | -1.0927
0.78135 | -33.9372 | -1.1131
0.78745 | -33.9387 | -1.1334
0.79355 | -33.94 | -1.1535
0.79966 | -33.9419 | -1.1732
0.80576 | -33.9443 | -1.1932
0.81187 | -33.9461 | -1.2152
0.81797 | -33.9482 | -1.2369
0.82408 | -33.95 | -1.2578
0.83018 | -33.9519 | -1.2787
0.83628 | -33.9539 | -1.2974
0.84239 | -33.9552 | -1.3184
0.84849 | -33.9567 | -1.3406
0.8546 | -33.9583 | -1.3607
0.8607 | -33.9596 | -1.3818
0.86681 | -33.9608 | -1.4011
0.87291 | -33.9619 | -1.4202
0.87901 | -33.9624 | -1.4386
0.88512 | -33.9632 | -1.4569
0.89122 | -33.964 | -1.4755
0.89733 | -33.9644 | -1.4912
0.90343 | -33.9649 | -1.509
0.90953 | -33.9651 | -1.5234
0.91564 | -33.9652 | -1.5393
0.92174 | -33.9655 | -1.5543
0.92785 | -33.9656 | -1.5689
0.93395 | -33.9657 | -1.5827
0.94006 | -33.9659 | -1.5955
0.94616 | -33.9656 | -1.6071
0.95226 | -33.9658 | -1.6196
0.95837 | -33.9659 | -1.6334
0.96447 | -33.9662 | -1.6461
0.97058 | -33.966 | -1.6593
0.97668 | -33.9658 | -1.6726
0.98279 | -33.9655 | -1.6852
0.98889 | -33.9654 | -1.6956
0.99499 | -33.9653 | -1.7089
1.0011 | -33.9652 | -1.7193
1.0072 | -33.9649 | -1.7307
1.0133 | -33.9645 | -1.7411
1.0194 | -33.9641 | -1.754
1.0255 | -33.9638 | -1.7676
1.0316 | -33.9632 | -1.7793
1.0377 | -33.9628 | -1.7929
1.0438 | -33.9621 | -1.8055
1.0499 | -33.9613 | -1.8208
1.056 | -33.9606 | -1.8351
1.0621 | -33.9599 | -1.8494
1.0682 | -33.9594 | -1.8617
1.0743 | -33.9587 | -1.8772
1.0805 | -33.9582 | -1.8916
1.0866 | -33.9572 | -1.9067
1.0927 | -33.9563 | -1.9212
1.0988 | -33.9556 | -1.9362
1.1049 | -33.9547 | -1.9504
1.111 | -33.9539 | -1.9658
1.1171 | -33.9531 | -1.9804
1.1232 | -33.9526 | -1.9957
1.1293 | -33.9522 | -2.0097
1.1354 | -33.9519 | -2.0248
1.1415 | -33.9514 | -2.0404
1.1476 | -33.9508 | -2.056
1.1537 | -33.9502 | -2.0708
1.1598 | -33.9503 | -2.0856
1.1659 | -33.9501 | -2.0989
1.172 | -33.9498 | -2.1128
1.1781 | -33.9495 | -2.1267
1.1842 | -33.9494 | -2.141
1.1903 | -33.9498 | -2.1518
1.1964 | -33.9497 | -2.1651
1.2025 | -33.9498 | -2.1791
1.2086 | -33.9501 | -2.1914
1.2147 | -33.9502 | -2.2038
1.2209 | -33.9505 | -2.2149
1.227 | -33.9514 | -2.2254
1.2331 | -33.9518 | -2.2373
1.2392 | -33.9522 | -2.249
1.2453 | -33.9527 | -2.2599
1.2514 | -33.9531 | -2.2695
1.2575 | -33.9538 | -2.2803
1.2636 | -33.9547 | -2.289
1.2697 | -33.9554 | -2.2979
1.2758 | -33.9561 | -2.3089
1.2819 | -33.9569 | -2.3186
1.288 | -33.9576 | -2.3292
1.2941 | -33.9586 | -2.3373
1.3002 | -33.9593 | -2.345
1.3063 | -33.9599 | -2.3537
1.3124 | -33.9606 | -2.3624
1.3185 | -33.9615 | -2.3712
1.3246 | -33.9622 | -2.38
1.3307 | -33.9632 | -2.3883
1.3368 | -33.964 | -2.3967
1.3429 | -33.9647 | -2.4053
1.349 | -33.9653 | -2.4144
1.3551 | -33.9661 | -2.4235
1.3613 | -33.9667 | -2.4314
1.3674 | -33.9674 | -2.4413
1.3735 | -33.9679 | -2.4508
1.3796 | -33.9684 | -2.4595
1.3857 | -33.9688 | -2.4685
1.3918 | -33.9693 | -2.4785
1.3979 | -33.97 | -2.4873
1.404 | -33.9705 | -2.4986
1.4101 | -33.971 | -2.5088
1.4162 | -33.9714 | -2.5192
1.4223 | -33.9719 | -2.5304
1.4284 | -33.9725 | -2.5402
1.4345 | -33.9727 | -2.5516
1.4406 | -33.9731 | -2.5652
1.4467 | -33.9736 | -2.5767
1.4528 | -33.974 | -2.5879
1.4589 | -33.9746 | -2.5997
1.465 | -33.9748 | -2.6104
1.4711 | -33.9754 | -2.6216
1.4772 | -33.9759 | -2.635
1.4833 | -33.9764 | -2.6481
1.4894 | -33.977 | -2.6602
1.4955 | -33.9776 | -2.6721
1.5016 | -33.9782 | -2.6846
1.5078 | -33.9786 | -2.6985
1.5139 | -33.979 | -2.7116
1.52 | -33.9795 | -2.7242
1.5261 | -33.9798 | -2.7367
1.5322 | -33.9801 | -2.7508
1.5383 | -33.9806 | -2.7634
1.5444 | -33.9811 | -2.776
1.5505 | -33.9816 | -2.7895
1.5566 | -33.9822 | -2.8025
1.5627 | -33.9827 | -2.8163
1.5688 | -33.9834 | -2.8299
1.5749 | -33.9837 | -2.8425
1.581 | -33.984 | -2.8559
1.5871 | -33.9845 | -2.8697
1.5932 | -33.9848 | -2.8825
1.5993 | -33.9853 | -2.8968
1.6054 | -33.9857 | -2.9101
1.6115 | -33.9861 | -2.9243
1.6176 | -33.9864 | -2.9368
1.6237 | -33.9867 | -2.9505
1.6298 | -33.9869 | -2.963
1.6359 | -33.9873 | -2.9757
1.642 | -33.9878 | -2.9875
1.6482 | -33.988 | -3.0005
1.6543 | -33.9884 | -3.0142
1.6604 | -33.9887 | -3.0262
1.6665 | -33.9888 | -3.0386
1.6726 | -33.989 | -3.051
1.6787 | -33.9894 | -3.0635
1.6848 | -33.9893 | -3.076
1.6909 | -33.9896 | -3.0887
1.697 | -33.9898 | -3.1008
1.7031 | -33.9899 | -3.1134
1.7092 | -33.9898 | -3.1254
1.7153 | -33.99 | -3.1373
1.7214 | -33.9898 | -3.1495
1.7275 | -33.9897 | -3.162
1.7336 | -33.9897 | -3.175
1.7397 | -33.9898 | -3.1872
1.7458 | -33.9898 | -3.1993
1.7519 | -33.9898 | -3.2114
1.758 | -33.9899 | -3.2233
1.7641 | -33.99 | -3.2348
1.7702 | -33.99 | -3.2477
1.7763 | -33.9899 | -3.2591
1.7824 | -33.9899 | -3.2715
1.7885 | -33.9899 | -3.284
1.7947 | -33.99 | -3.2959
1.8008 | -33.99 | -3.307
1.8069 | -33.99 | -3.3185
1.813 | -33.9901 | -3.3306
1.8191 | -33.9901 | -3.3434
1.8252 | -33.9901 | -3.3562
1.8313 | -33.9903 | -3.3684
1.8374 | -33.9904 | -3.3802
1.8435 | -33.9905 | -3.3921
1.8496 | -33.9906 | -3.4052
1.8557 | -33.9908 | -3.418
1.8618 | -33.991 | -3.4303
1.8679 | -33.9911 | -3.4423
1.874 | -33.9913 | -3.4543
1.8801 | -33.9914 | -3.4667
1.8862 | -33.9917 | -3.4787
1.8923 | -33.9918 | -3.4909
1.8984 | -33.992 | -3.503
1.9045 | -33.992 | -3.5146
1.9106 | -33.9921 | -3.5268
1.9167 | -33.9921 | -3.5391
1.9228 | -33.9923 | -3.5509
1.9289 | -33.9924 | -3.5629
1.9351 | -33.9926 | -3.5737
1.9412 | -33.9928 | -3.5849
1.9473 | -33.9931 | -3.597
1.9534 | -33.9933 | -3.6085
1.9595 | -33.9935 | -3.6203
1.9656 | -33.9938 | -3.6328
1.9717 | -33.9941 | -3.6445
1.9778 | -33.9944 | -3.6559
1.9839 | -33.9946 | -3.6675
1.99 | -33.9948 | -3.6785
1.9961 | -33.9949 | -3.6908
2.0022 | -33.9952 | -3.7023
2.0083 | -33.9953 | -3.7135
2.0144 | -33.9955 | -3.7249
2.0205 | -33.9958 | -3.7367
2.0266 | -33.9962 | -3.7481
2.0327 | -33.9964 | -3.7598
2.0388 | -33.9969 | -3.7715
2.0449 | -33.9971 | -3.7829
2.051 | -33.9974 | -3.7943
2.0571 | -33.9977 | -3.8057
2.0632 | -33.9979 | -3.8172
2.0693 | -33.9981 | -3.8295
2.0754 | -33.9984 | -3.8408
2.0816 | -33.9988 | -3.8531
2.0877 | -33.999 | -3.8647
2.0938 | -33.9994 | -3.8769
2.0999 | -33.9995 | -3.8886
2.106 | -33.9998 | -3.9003
2.1121 | -34 | -3.9124
2.1182 | -34.0002 | -3.9231
2.1243 | -34.0005 | -3.9339
2.1304 | -34.0008 | -3.9451
2.1365 | -34.0011 | -3.9569
2.1426 | -34.0013 | -3.9689
2.1487 | -34.0016 | -3.9802
2.1548 | -34.0021 | -3.9921
2.1609 | -34.0025 | -4.0036
2.167 | -34.003 | -4.0148
2.1731 | -34.0033 | -4.0269
2.1792 | -34.0035 | -4.038
2.1853 | -34.004 | -4.05
2.1914 | -34.0043 | -4.0614
2.1975 | -34.0047 | -4.073
2.2036 | -34.0049 | -4.0852
2.2097 | -34.0052 | -4.0967
2.2158 | -34.0056 | -4.1073
2.222 | -34.0059 | -4.1178
2.2281 | -34.0062 | -4.1292
2.2342 | -34.0065 | -4.1406
2.2403 | -34.0067 | -4.1513
2.2464 | -34.007 | -4.162
2.2525 | -34.0073 | -4.1734
2.2586 | -34.0077 | -4.184
2.2647 | -34.008 | -4.1944
2.2708 | -34.0083 | -4.2056
2.2769 | -34.0087 | -4.2163
2.283 | -34.0089 | -4.2275
2.2891 | -34.0092 | -4.2384
2.2952 | -34.0094 | -4.2489
2.3013 | -34.0096 | -4.2598
2.3074 | -34.0099 | -4.2702
2.3135 | -34.0101 | -4.2816
2.3196 | -34.0104 | -4.2917
2.3257 | -34.0106 | -4.3024
2.3318 | -34.0108 | -4.3135
2.3379 | -34.011 | -4.3243
2.344 | -34.0111 | -4.3357
2.3501 | -34.0114 | -4.3464
2.3562 | -34.0115 | -4.3567
2.3623 | -34.0117 | -4.367
2.3685 | -34.0118 | -4.3775
2.3746 | -34.012 | -4.3887
2.3807 | -34.0122 | -4.3992
2.3868 | -34.0124 | -4.4099
2.3929 | -34.0126 | -4.4205
2.399 | -34.0127 | -4.4309
2.4051 | -34.013 | -4.4414
2.4112 | -34.0132 | -4.452
2.4173 | -34.0134 | -4.4626
2.4234 | -34.0135 | -4.4734
2.4295 | -34.0137 | -4.4842
2.4356 | -34.0138 | -4.495
2.4417 | -34.014 | -4.5061
2.4478 | -34.0141 | -4.517
2.4539 | -34.0143 | -4.5271
2.46 | -34.0144 | -4.5386
2.4661 | -34.0146 | -4.5496
2.4722 | -34.0148 | -4.5597
2.4783 | -34.0149 | -4.5701
2.4844 | -34.015 | -4.5811
2.4905 | -34.0151 | -4.592
2.4966 | -34.0152 | -4.6027
2.5027 | -34.0153 | -4.6132
2.5089 | -34.0155 | -4.624
2.515 | -34.0155 | -4.6347
2.5211 | -34.0156 | -4.6454
2.5272 | -34.0157 | -4.6563
2.5333 | -34.0158 | -4.6672
2.5394 | -34.016 | -4.6779
2.5455 | -34.0161 | -4.6893
2.5516 | -34.0163 | -4.6995
2.5577 | -34.0165 | -4.7109
2.5638 | -34.0167 | -4.7226
2.5699 | -34.0167 | -4.734
2.576 | -34.0168 | -4.7456
2.5821 | -34.0169 | -4.7568
2.5882 | -34.017 | -4.7678
2.5943 | -34.0171 | -4.7786
2.6004 | -34.0173 | -4.7895
2.6065 | -34.0174 | -4.7998
2.6126 | -34.0176 | -4.8102
2.6187 | -34.0177 | -4.8205
2.6248 | -34.0178 | -4.8312
2.6309 | -34.018 | -4.8426
2.637 | -34.0183 | -4.8542
2.6431 | -34.0184 | -4.8646
2.6492 | -34.0185 | -4.8759
2.6554 | -34.0187 | -4.8867
2.6615 | -34.0188 | -4.8976
2.6676 | -34.0191 | -4.909
2.6737 | -34.0194 | -4.9201
2.6798 | -34.0197 | -4.9311
2.6859 | -34.0199 | -4.9423
2.692 | -34.02 | -4.9531
2.6981 | -34.0202 | -4.9642
2.7042 | -34.0204 | -4.9753
2.7103 | -34.0206 | -4.9858
2.7164 | -34.021 | -4.9967
2.7225 | -34.0212 | -5.0081
2.7286 | -34.0214 | -5.019
2.7347 | -34.0217 | -5.0294
2.7408 | -34.0219 | -5.0394
2.7469 | -34.0222 | -5.0507
2.753 | -34.0225 | -5.0612
2.7591 | -34.0228 | -5.0724
2.7652 | -34.023 | -5.0839
2.7713 | -34.0232 | -5.0945
2.7774 | -34.0234 | -5.1054
2.7835 | -34.0236 | -5.1163
2.7896 | -34.0238 | -5.1268
2.7958 | -34.0241 | -5.1375
2.8019 | -34.0244 | -5.1478
2.808 | -34.0247 | -5.1585
2.8141 | -34.025 | -5.1696
2.8202 | -34.0253 | -5.1804
2.8263 | -34.0257 | -5.1912
2.8324 | -34.026 | -5.2026
2.8385 | -34.0264 | -5.2134
2.8446 | -34.0268 | -5.224
2.8507 | -34.0272 | -5.2348
2.8568 | -34.0276 | -5.2452
2.8629 | -34.0278 | -5.2557
2.869 | -34.0282 | -5.2668
2.8751 | -34.0285 | -5.2775
2.8812 | -34.0289 | -5.2888
2.8873 | -34.0292 | -5.2996
2.8934 | -34.0296 | -5.3106
2.8995 | -34.0299 | -5.3217
2.9056 | -34.0302 | -5.3328
2.9117 | -34.0306 | -5.3438
2.9178 | -34.0309 | -5.3545
2.9239 | -34.0314 | -5.366
2.93 | -34.0318 | -5.3775
2.9361 | -34.0321 | -5.3883
2.9423 | -34.0324 | -5.3989
2.9484 | -34.0329 | -5.41
2.9545 | -34.0333 | -5.4209
2.9606 | -34.0337 | -5.4323
2.9667 | -34.0341 | -5.4435
2.9728 | -34.0344 | -5.4548
2.9789 | -34.0349 | -5.4658
2.985 | -34.0354 | -5.4772
2.9911 | -34.0357 | -5.4885
2.9972 | -34.0362 | -5.4995
3.0033 | -34.0367 | -5.5111
3.0094 | -34.0371 | -5.5221
3.0155 | -34.0375 | -5.5331
3.0216 | -34.0379 | -5.5443
3.0277 | -34.0382 | -5.556
3.0338 | -34.0387 | -5.5675
3.0399 | -34.0391 | -5.5788
3.046 | -34.0394 | -5.5904
3.0521 | -34.0398 | -5.6015
3.0582 | -34.0401 | -5.6136
3.0643 | -34.0405 | -5.6251
3.0704 | -34.0408 | -5.637
3.0765 | -34.0411 | -5.6481
3.0827 | -34.0414 | -5.6594
3.0888 | -34.0418 | -5.6707
3.0949 | -34.0422 | -5.6821
3.101 | -34.0426 | -5.6936
3.1071 | -34.0429 | -5.7054
3.1132 | -34.0433 | -5.7172
3.1193 | -34.0437 | -5.7288
3.1254 | -34.0441 | -5.7401
3.1315 | -34.0445 | -5.7521
3.1376 | -34.0448 | -5.7636
3.1437 | -34.0452 | -5.7752
3.1498 | -34.0456 | -5.7873
3.1559 | -34.0459 | -5.7988
3.162 | -34.0462 | -5.811
3.1681 | -34.0465 | -5.823
3.1742 | -34.0468 | -5.8347
3.1803 | -34.0471 | -5.8464
3.1864 | -34.0475 | -5.8577
3.1925 | -34.0479 | -5.8693
3.1986 | -34.0481 | -5.8814
3.2047 | -34.0484 | -5.8935
3.2108 | -34.0487 | -5.9047
3.2169 | -34.0491 | -5.9165
3.223 | -34.0493 | -5.9283
3.2292 | -34.0496 | -5.9402
3.2353 | -34.0498 | -5.951
3.2414 | -34.05 | -5.9625
3.2475 | -34.0503 | -5.9739
3.2536 | -34.0506 | -5.9854
3.2597 | -34.0509 | -5.9964
3.2658 | -34.0511 | -6.0074
3.2719 | -34.0513 | -6.0184
3.278 | -34.0516 | -6.0301
3.2841 | -34.0519 | -6.0412
3.2902 | -34.0522 | -6.0527
3.2963 | -34.0524 | -6.0642
3.3024 | -34.0525 | -6.0753
3.3085 | -34.0529 | -6.087
3.3146 | -34.053 | -6.0982
3.3207 | -34.0533 | -6.1097
3.3268 | -34.0535 | -6.1212
3.3329 | -34.0537 | -6.132
3.339 | -34.0539 | -6.1428
3.3451 | -34.0541 | -6.1541
3.3512 | -34.0543 | -6.1648
3.3573 | -34.0546 | -6.1761
3.3634 | -34.0547 | -6.1872
3.3696 | -34.055 | -6.1986
3.3757 | -34.0551 | -6.2094
3.3818 | -34.0554 | -6.2203
3.3879 | -34.0555 | -6.2314
3.394 | -34.0557 | -6.2428
3.4001 | -34.0559 | -6.2536
3.4062 | -34.056 | -6.2646
3.4123 | -34.0562 | -6.2752
3.4184 | -34.0564 | -6.2865
3.4245 | -34.0566 | -6.2973
3.4306 | -34.0567 | -6.308
3.4367 | -34.0569 | -6.3185
3.4428 | -34.0571 | -6.3289
3.4489 | -34.0572 | -6.3398
3.455 | -34.0574 | -6.3507
3.4611 | -34.0576 | -6.362
3.4672 | -34.0577 | -6.373
3.4733 | -34.0579 | -6.3841
3.4794 | -34.0581 | -6.3949
3.4855 | -34.0582 | -6.4053
3.4916 | -34.0584 | -6.4157
3.4977 | -34.0586 | -6.427
3.5038 | -34.0587 | -6.4379
3.5099 | -34.0589 | -6.4481
3.5161 | -34.059 | -6.4584
3.5222 | -34.0592 | -6.4688
3.5283 | -34.0592 | -6.4795
3.5344 | -34.0594 | -6.4903
3.5405 | -34.0596 | -6.5009
3.5466 | -34.0598 | -6.5118
3.5527 | -34.0598 | -6.5226
3.5588 | -34.06 | -6.5332
3.5649 | -34.06 | -6.5438
3.571 | -34.0602 | -6.5544
3.5771 | -34.0603 | -6.5655
3.5832 | -34.0604 | -6.5754
3.5893 | -34.0607 | -6.586
3.5954 | -34.0608 | -6.5968
3.6015 | -34.061 | -6.6077
3.6076 | -34.0612 | -6.6182
3.6137 | -34.0614 | -6.6283
3.6198 | -34.0617 | -6.6385
3.6259 | -34.0618 | -6.6494
3.632 | -34.062 | -6.66
3.6381 | -34.0621 | -6.6708
3.6442 | -34.0623 | -6.6813
3.6503 | -34.0625 | -6.6915
3.6565 | -34.0627 | -6.7022
3.6626 | -34.0628 | -6.7128
3.6687 | -34.0631 | -6.7235
3.6748 | -34.0633 | -6.7339
3.6809 | -34.0636 | -6.7445
3.687 | -34.0638 | -6.7555
3.6931 | -34.0641 | -6.7657
3.6992 | -34.0643 | -6.7763
3.7053 | -34.0645 | -6.7868
3.7114 | -34.0646 | -6.7975
3.7175 | -34.0648 | -6.8082
3.7236 | -34.0651 | -6.8182
3.7297 | -34.0653 | -6.829
3.7358 | -34.0655 | -6.8399
3.7419 | -34.0657 | -6.8507
3.748 | -34.066 | -6.8613
3.7541 | -34.0662 | -6.8719
3.7602 | -34.0665 | -6.8823
3.7663 | -34.0667 | -6.8928
3.7724 | -34.0669 | -6.9035
3.7785 | -34.0672 | -6.9138
3.7846 | -34.0674 | -6.9242
3.7907 | -34.0677 | -6.9352
3.7969 | -34.068 | -6.946
3.803 | -34.0683 | -6.9561
3.8091 | -34.0685 | -6.9667
3.8152 | -34.0688 | -6.9772
3.8213 | -34.069 | -6.988
3.8274 | -34.0693 | -6.9983
3.8335 | -34.0696 | -7.0091
3.8396 | -34.07 | -7.0197
3.8457 | -34.0703 | -7.0305
3.8518 | -34.0706 | -7.041
3.8579 | -34.071 | -7.0515
3.864 | -34.0713 | -7.0623
3.8701 | -34.0716 | -7.0734
3.8762 | -34.0718 | -7.0838
3.8823 | -34.0721 | -7.0946
3.8884 | -34.0724 | -7.1053
3.8945 | -34.0728 | -7.1162
3.9006 | -34.0731 | -7.1273
3.9067 | -34.0734 | -7.1384
3.9128 | -34.0737 | -7.1491
3.9189 | -34.074 | -7.1596
3.925 | -34.0743 | -7.1704
3.9311 | -34.0746 | -7.1806
3.9372 | -34.0749 | -7.1915
3.9434 | -34.0752 | -7.2022
3.9495 | -34.0756 | -7.2131
3.9556 | -34.076 | -7.2235
3.9617 | -34.0763 | -7.2338
3.9678 | -34.0766 | -7.2442
3.9739 | -34.0769 | -7.2549
3.98 | -34.0772 | -7.2657
3.9861 | -34.0775 | -7.2763
3.9922 | -34.0777 | -7.287
3.9983 | -34.0781 | -7.2973
4.0044 | -34.0784 | -7.3075
4.0105 | -34.0787 | -7.3182
4.0166 | -34.079 | -7.3289
4.0227 | -34.0793 | -7.3395
4.0288 | -34.0796 | -7.3507
4.0349 | -34.0798 | -7.3611
4.041 | -34.0801 | -7.3717
4.0471 | -34.0804 | -7.3825
4.0532 | -34.0807 | -7.3927
4.0593 | -34.081 | -7.4036
4.0654 | -34.0813 | -7.414
4.0715 | -34.0816 | -7.4251
4.0776 | -34.082 | -7.4357
4.0838 | -34.0823 | -7.4468
4.0899 | -34.0826 | -7.4572
4.096 | -34.0829 | -7.4677
4.1021 | -34.0832 | -7.4781
4.1082 | -34.0834 | -7.4885
4.1143 | -34.0837 | -7.4988
4.1204 | -34.084 | -7.5091
4.1265 | -34.0843 | -7.5193
4.1326 | -34.0846 | -7.5296
4.1387 | -34.085 | -7.54
4.1448 | -34.0853 | -7.5507
4.1509 | -34.0855 | -7.5612
4.157 | -34.0858 | -7.5723
4.1631 | -34.0861 | -7.5831
4.1692 | -34.0864 | -7.5934
4.1753 | -34.0867 | -7.6039
4.1814 | -34.0869 | -7.6141
4.1875 | -34.0872 | -7.6249
4.1936 | -34.0875 | -7.6356
4.1997 | -34.0877 | -7.6458
4.2058 | -34.088 | -7.6566
4.2119 | -34.0882 | -7.6669
4.218 | -34.0885 | -7.6771
4.2241 | -34.0887 | -7.6874
4.2303 | -34.0891 | -7.6977
4.2364 | -34.0893 | -7.7082
4.2425 | -34.0895 | -7.7187
4.2486 | -34.0897 | -7.7293
4.2547 | -34.0899 | -7.7395
4.2608 | -34.0901 | -7.7501
4.2669 | -34.0903 | -7.7602
4.273 | -34.0906 | -7.7708
4.2791 | -34.0908 | -7.7812
4.2852 | -34.0911 | -7.7918
4.2913 | -34.0912 | -7.8019
4.2974 | -34.0915 | -7.8126
4.3035 | -34.0918 | -7.8234
4.3096 | -34.092 | -7.834
4.3157 | -34.0922 | -7.8442
4.3218 | -34.0924 | -7.8545
4.3279 | -34.0927 | -7.865
4.334 | -34.0929 | -7.8753
4.3401 | -34.0931 | -7.8852
4.3462 | -34.0934 | -7.8957
4.3523 | -34.0937 | -7.9056
4.3584 | -34.0938 | -7.9156
4.3645 | -34.0941 | -7.9258
4.3707 | -34.0943 | -7.9359
4.3768 | -34.0945 | -7.9461
4.3829 | -34.0948 | -7.9569
4.389 | -34.095 | -7.9672
4.3951 | -34.0952 | -7.9773
4.4012 | -34.0954 | -7.9879
4.4073 | -34.0956 | -7.9981
4.4134 | -34.0958 | -8.0083
4.4195 | -34.0961 | -8.0185
4.4256 | -34.0964 | -8.0286
4.4317 | -34.0966 | -8.0388
4.4378 | -34.0968 | -8.0487
4.4439 | -34.097 | -8.0586
4.45 | -34.0971 | -8.0689
4.4561 | -34.0973 | -8.0791
4.4622 | -34.0975 | -8.0894
4.4683 | -34.0977 | -8.0996
4.4744 | -34.0979 | -8.1102
4.4805 | -34.0982 | -8.12
4.4866 | -34.0984 | -8.1302
4.4927 | -34.0986 | -8.1407
4.4988 | -34.0988 | -8.1509
4.5049 | -34.0991 | -8.1607
4.511 | -34.0993 | -8.1706
4.5172 | -34.0995 | -8.1807
4.5233 | -34.0997 | -8.1913
4.5294 | -34.1 | -8.2013
4.5355 | -34.1002 | -8.2116
4.5416 | -34.1004 | -8.2217
4.5477 | -34.1006 | -8.2319
4.5538 | -34.1008 | -8.2416
4.5599 | -34.1011 | -8.2517
4.566 | -34.1013 | -8.2619
4.5721 | -34.1015 | -8.2721
4.5782 | -34.1017 | -8.2825
4.5843 | -34.102 | -8.2928
4.5904 | -34.1022 | -8.303
4.5965 | -34.1025 | -8.3131
4.6026 | -34.1027 | -8.3237
4.6087 | -34.1029 | -8.3338
4.6148 | -34.1031 | -8.3443
4.6209 | -34.1033 | -8.3544
4.627 | -34.1036 | -8.3648
4.6331 | -34.1037 | -8.3748
4.6392 | -34.104 | -8.3851
4.6453 | -34.1042 | -8.3953
4.6514 | -34.1045 | -8.4055
4.6576 | -34.1047 | -8.4161
4.6637 | -34.1048 | -8.4259
4.6698 | -34.105 | -8.4359
4.6759 | -34.1053 | -8.4465
4.682 | -34.1055 | -8.4566
4.6881 | -34.1056 | -8.4667
4.6942 | -34.1059 | -8.4769
4.7003 | -34.1061 | -8.4871
4.7064 | -34.1063 | -8.4971
4.7125 | -34.1066 | -8.5076
4.7186 | -34.1067 | -8.5175
4.7247 | -34.107 | -8.5278
4.7308 | -34.1073 | -8.5379
4.7369 | -34.1075 | -8.5482
4.743 | -34.1077 | -8.5588
4.7491 | -34.1079 | -8.5693
4.7552 | -34.1081 | -8.5797
4.7613 | -34.1083 | -8.59
4.7674 | -34.1085 | -8.6
4.7735 | -34.1087 | -8.6107
4.7796 | -34.109 | -8.6213
4.7857 | -34.1092 | -8.6316
4.7918 | -34.1095 | -8.6417
4.7979 | -34.1098 | -8.6518
4.8041 | -34.11 | -8.6621
4.8102 | -34.1102 | -8.672
4.8163 | -34.1104 | -8.6822
4.8224 | -34.1107 | -8.6924
4.8285 | -34.1109 | -8.7024
4.8346 | -34.1112 | -8.7124
4.8407 | -34.1114 | -8.7227
4.8468 | -34.1116 | -8.7332
4.8529 | -34.1119 | -8.7438
4.859 | -34.1121 | -8.7544
4.8651 | -34.1124 | -8.7644
4.8712 | -34.1126 | -8.7744
4.8773 | -34.1128 | -8.7848
4.8834 | -34.1131 | -8.795
4.8895 | -34.1133 | -8.8051
4.8956 | -34.1135 | -8.8153
4.9017 | -34.1137 | -8.8254
4.9078 | -34.1139 | -8.836
4.9139 | -34.1142 | -8.8458
4.92 | -34.1145 | -8.8557
4.9261 | -34.1148 | -8.8656
4.9322 | -34.115 | -8.8759
4.9383 | -34.1153 | -8.8861
4.9445 | -34.1155 | -8.8964
4.9506 | -34.1158 | -8.9064
4.9567 | -34.116 | -8.9164
4.9628 | -34.1162 | -8.9266
4.9689 | -34.1165 | -8.9373
4.975 | -34.1167 | -8.9478
4.9811 | -34.117 | -8.9578
4.9872 | -34.1173 | -8.9679
4.9933 | -34.1176 | -8.9784
4.9994 | -34.1179 | -8.9891
5.0055 | -34.1181 | -8.9994
5.0116 | -34.1184 | -9.0102
5.0177 | -34.1187 | -9.0207
5.0238 | -34.119 | -9.0313
5.0299 | -34.1193 | -9.042
5.036 | -34.1196 | -9.0523
5.0421 | -34.1198 | -9.0625
5.0482 | -34.1201 | -9.0728
5.0543 | -34.1204 | -9.0829
5.0604 | -34.1207 | -9.0931
5.0665 | -34.121 | -9.1032
5.0726 | -34.1213 | -9.1138
5.0787 | -34.1216 | -9.1243
5.0848 | -34.1218 | -9.1347
5.091 | -34.1221 | -9.1448
5.0971 | -34.1224 | -9.1548
5.1032 | -34.1227 | -9.1655
5.1093 | -34.123 | -9.1758
5.1154 | -34.1232 | -9.1858
5.1215 | -34.1235 | -9.1959
5.1276 | -34.1238 | -9.206
5.1337 | -34.1241 | -9.216
5.1398 | -34.1243 | -9.226
5.1459 | -34.1246 | -9.2366
5.152 | -34.1249 | -9.2471
5.1581 | -34.1252 | -9.2576
5.1642 | -34.1255 | -9.268
5.1703 | -34.1258 | -9.2783
5.1764 | -34.1261 | -9.2885
5.1825 | -34.1264 | -9.2985
5.1886 | -34.1267 | -9.3085
5.1947 | -34.1269 | -9.3188
5.2008 | -34.1273 | -9.3293
5.2069 | -34.1275 | -9.3399
5.213 | -34.1278 | -9.3499
5.2191 | -34.1282 | -9.3603
5.2252 | -34.1284 | -9.3707
5.2314 | -34.1287 | -9.3812
5.2375 | -34.129 | -9.3913
5.2436 | -34.1292 | -9.4017
5.2497 | -34.1296 | -9.4117
5.2558 | -34.1299 | -9.4224
5.2619 | -34.1302 | -9.4331
5.268 | -34.1305 | -9.4436
5.2741 | -34.1308 | -9.454
5.2802 | -34.1311 | -9.4648
5.2863 | -34.1315 | -9.475
5.2924 | -34.1318 | -9.4856
5.2985 | -34.1321 | -9.496
5.3046 | -34.1324 | -9.5065
5.3107 | -34.1327 | -9.5171
5.3168 | -34.133 | -9.5275
5.3229 | -34.1333 | -9.5379
5.329 | -34.1336 | -9.5483
5.3351 | -34.1339 | -9.5584
5.3412 | -34.1342 | -9.5687
5.3473 | -34.1345 | -9.5792
5.3534 | -34.1348 | -9.5897
5.3595 | -34.1351 | -9.6002
5.3656 | -34.1355 | -9.6105
5.3717 | -34.1357 | -9.6213
5.3779 | -34.136 | -9.6321
5.384 | -34.1363 | -9.6425
5.3901 | -34.1367 | -9.6532
5.3962 | -34.137 | -9.6634
5.4023 | -34.1373 | -9.674
5.4084 | -34.1377 | -9.6844
5.4145 | -34.138 | -9.6946
5.4206 | -34.1383 | -9.705
5.4267 | -34.1386 | -9.7158
5.4328 | -34.1389 | -9.7263
5.4389 | -34.1392 | -9.7365
5.445 | -34.1395 | -9.7473
5.4511 | -34.1398 | -9.7575
5.4572 | -34.1402 | -9.768
5.4633 | -34.1405 | -9.7785
5.4694 | -34.1407 | -9.7889
5.4755 | -34.141 | -9.7991
5.4816 | -34.1414 | -9.8096
5.4877 | -34.1417 | -9.8199
5.4938 | -34.142 | -9.83
5.4999 | -34.1423 | -9.8404
5.506 | -34.1427 | -9.8508
5.5121 | -34.143 | -9.8611
5.5183 | -34.1433 | -9.8714
5.5244 | -34.1436 | -9.882
5.5305 | -34.1439 | -9.8924
5.5366 | -34.1442 | -9.9026
5.5427 | -34.1445 | -9.9131
5.5488 | -34.1448 | -9.9236
5.5549 | -34.1451 | -9.9341
5.561 | -34.1454 | -9.9448
5.5671 | -34.1458 | -9.9551
5.5732 | -34.1461 | -9.9658
5.5793 | -34.1464 | -9.9766
5.5854 | -34.1468 | -9.987
5.5915 | -34.147 | -9.9977
5.5976 | -34.1473 | -10.0083
5.6037 | -34.1476 | -10.0188
5.6098 | -34.1479 | -10.0292
5.6159 | -34.1483 | -10.0395
5.622 | -34.1486 | -10.05
5.6281 | -34.1489 | -10.0606
5.6342 | -34.1492 | -10.0714
5.6403 | -34.1495 | -10.0819
5.6464 | -34.1498 | -10.0925
5.6525 | -34.1501 | -10.1031
5.6586 | -34.1505 | -10.1132
5.6648 | -34.1508 | -10.1237
5.6709 | -34.1512 | -10.1344
5.677 | -34.1515 | -10.1449
5.6831 | -34.1518 | -10.1555
5.6892 | -34.1521 | -10.1661
5.6953 | -34.1524 | -10.1766
5.7014 | -34.1528 | -10.1872
5.7075 | -34.153 | -10.1977
5.7136 | -34.1533 | -10.2081
5.7197 | -34.1536 | -10.2185
5.7258 | -34.1539 | -10.2288
5.7319 | -34.1543 | -10.2394
5.738 | -34.1546 | -10.2498
5.7441 | -34.155 | -10.2603
5.7502 | -34.1553 | -10.2704
5.7563 | -34.1556 | -10.2809
5.7624 | -34.156 | -10.2916
5.7685 | -34.1563 | -10.3021
5.7746 | -34.1566 | -10.3129
5.7807 | -34.1568 | -10.3234
5.7868 | -34.1571 | -10.3341
5.7929 | -34.1574 | -10.3445
5.799 | -34.1577 | -10.3549
5.8052 | -34.158 | -10.3652
5.8113 | -34.1584 | -10.3756
5.8174 | -34.1586 | -10.3859
5.8235 | -34.159 | -10.3963
5.8296 | -34.1593 | -10.407
5.8357 | -34.1595 | -10.4174
5.8418 | -34.1599 | -10.428
5.8479 | -34.1601 | -10.4385
5.854 | -34.1605 | -10.4492
5.8601 | -34.1608 | -10.4595
5.8662 | -34.1611 | -10.4701
5.8723 | -34.1614 | -10.4804
5.8784 | -34.1617 | -10.4908
5.8845 | -34.162 | -10.5011
5.8906 | -34.1623 | -10.5116
5.8967 | -34.1626 | -10.5219
5.9028 | -34.1629 | -10.5322
5.9089 | -34.1633 | -10.5425
5.915 | -34.1636 | -10.5527
5.9211 | -34.1639 | -10.5631
5.9272 | -34.1642 | -10.5732
5.9333 | -34.1645 | -10.5836
5.9394 | -34.1648 | -10.594
5.9455 | -34.1652 | -10.6042
5.9517 | -34.1655 | -10.6144
5.9578 | -34.1659 | -10.625
5.9639 | -34.1662 | -10.635
5.97 | -34.1665 | -10.6455
5.9761 | -34.1669 | -10.6558
5.9822 | -34.1672 | -10.6663
5.9883 | -34.1675 | -10.6766
5.9944 | -34.1678 | -10.6872
6.0005 | -34.1682 | -10.6977
6.0066 | -34.1685 | -10.7079
6.0127 | -34.1688 | -10.718
6.0188 | -34.1691 | -10.7287
6.0249 | -34.1694 | -10.7392
6.031 | -34.1697 | -10.7494
6.0371 | -34.17 | -10.7598
6.0432 | -34.1703 | -10.77
6.0493 | -34.1707 | -10.7802
6.0554 | -34.171 | -10.7904
6.0615 | -34.1714 | -10.8008
6.0676 | -34.1717 | -10.8114
6.0737 | -34.172 | -10.8219
6.0798 | -34.1723 | -10.8322
6.0859 | -34.1727 | -10.8424
6.0921 | -34.173 | -10.8525
6.0982 | -34.1733 | -10.8629
6.1043 | -34.1737 | -10.8731
6.1104 | -34.174 | -10.8837
6.1165 | -34.1744 | -10.8935
6.1226 | -34.1747 | -10.9038
6.1287 | -34.175 | -10.914
6.1348 | -34.1753 | -10.9243
6.1409 | -34.1757 | -10.9346
6.147 | -34.176 | -10.9448
6.1531 | -34.1763 | -10.9551
6.1592 | -34.1766 | -10.9652
6.1653 | -34.1769 | -10.9754
6.1714 | -34.1772 | -10.9857
6.1775 | -34.1776 | -10.9961
6.1836 | -34.1779 | -11.0065
6.1897 | -34.1782 | -11.0167
6.1958 | -34.1785 | -11.0273
6.2019 | -34.1789 | -11.0377
6.208 | -34.1792 | -11.048
6.2141 | -34.1795 | -11.0582
6.2202 | -34.1799 | -11.0686
6.2263 | -34.1802 | -11.0788
6.2325 | -34.1805 | -11.0887
6.2386 | -34.1808 | -11.0987
6.2447 | -34.1812 | -11.1092
6.2508 | -34.1815 | -11.1194
6.2569 | -34.1819 | -11.1296
6.263 | -34.1822 | -11.1399
6.2691 | -34.1825 | -11.15
6.2752 | -34.1828 | -11.1603
6.2813 | -34.1831 | -11.1706
6.2874 | -34.1834 | -11.1807
6.2935 | -34.1837 | -11.1908
6.2996 | -34.184 | -11.201
6.3057 | -34.1844 | -11.2113
6.3118 | -34.1846 | -11.2217
6.3179 | -34.185 | -11.2318
6.324 | -34.1853 | -11.242
6.3301 | -34.1856 | -11.2522
6.3362 | -34.186 | -11.2623
6.3423 | -34.1863 | -11.2727
6.3484 | -34.1866 | -11.283
6.3545 | -34.1869 | -11.2928
6.3606 | -34.1872 | -11.3028
6.3667 | -34.1875 | -11.3128
6.3728 | -34.1879 | -11.323
6.379 | -34.1882 | -11.3332
6.3851 | -34.1885 | -11.3432
6.3912 | -34.1888 | -11.3533
6.3973 | -34.1891 | -11.3635
6.4034 | -34.1894 | -11.3734
6.4095 | -34.1897 | -11.3836
6.4156 | -34.19 | -11.3938
6.4217 | -34.1903 | -11.4039
6.4278 | -34.1907 | -11.4142
6.4339 | -34.191 | -11.4242
6.44 | -34.1914 | -11.4343
6.4461 | -34.1916 | -11.4447
6.4522 | -34.192 | -11.4545
6.4583 | -34.1922 | -11.4649
6.4644 | -34.1925 | -11.4751
6.4705 | -34.1929 | -11.4849
6.4766 | -34.1932 | -11.495
6.4827 | -34.1935 | -11.5052
6.4888 | -34.1938 | -11.5152
6.4949 | -34.1941 | -11.525
6.501 | -34.1944 | -11.5352
6.5071 | -34.1947 | -11.5451
6.5132 | -34.1951 | -11.5552
6.5194 | -34.1954 | -11.5654
6.5255 | -34.1957 | -11.5753
6.5316 | -34.196 | -11.5855
6.5377 | -34.1963 | -11.5956
6.5438 | -34.1966 | -11.6054
6.5499 | -34.1969 | -11.6154
6.556 | -34.1972 | -11.6255
6.5621 | -34.1975 | -11.6358
6.5682 | -34.1978 | -11.6459
6.5743 | -34.1982 | -11.6559
6.5804 | -34.1985 | -11.6661
6.5865 | -34.1988 | -11.6762
6.5926 | -34.1991 | -11.6862
6.5987 | -34.1995 | -11.6959
6.6048 | -34.1998 | -11.706
6.6109 | -34.2001 | -11.7161
6.617 | -34.2004 | -11.7263
6.6231 | -34.2007 | -11.7363
6.6292 | -34.201 | -11.7463
6.6353 | -34.2013 | -11.7567
6.6414 | -34.2017 | -11.7667
6.6475 | -34.202 | -11.7767
6.6536 | -34.2023 | -11.7866
6.6597 | -34.2026 | -11.7963
6.6659 | -34.2029 | -11.8064
6.672 | -34.2032 | -11.8163
6.6781 | -34.2035 | -11.8263
6.6842 | -34.2039 | -11.8362
6.6903 | -34.2043 | -11.8462
6.6964 | -34.2046 | -11.8562
6.7025 | -34.2049 | -11.8661
6.7086 | -34.2052 | -11.8761
6.7147 | -34.2055 | -11.8862
6.7208 | -34.2058 | -11.8963
6.7269 | -34.2061 | -11.9062
6.733 | -34.2064 | -11.916
6.7391 | -34.2067 | -11.926
6.7452 | -34.2071 | -11.9357
6.7513 | -34.2074 | -11.9454
6.7574 | -34.2077 | -11.9555
6.7635 | -34.208 | -11.9653
6.7696 | -34.2083 | -11.9751
6.7757 | -34.2087 | -11.9851
6.7818 | -34.209 | -11.9951
6.7879 | -34.2094 | -12.0049
6.794 | -34.2097 | -12.0146
6.8001 | -34.21 | -12.0245
6.8063 | -34.2103 | -12.0345
6.8124 | -34.2106 | -12.0446
6.8185 | -34.2109 | -12.0547
6.8246 | -34.2112 | -12.0644
6.8307 | -34.2115 | -12.0742
6.8368 | -34.2118 | -12.0844
6.8429 | -34.2122 | -12.0944
6.849 | -34.2125 | -12.1044
6.8551 | -34.2129 | -12.1142
6.8612 | -34.2132 | -12.1241
6.8673 | -34.2135 | -12.134
6.8734 | -34.2138 | -12.144
6.8795 | -34.2141 | -12.154
6.8856 | -34.2144 | -12.1639
6.8917 | -34.2148 | -12.1738
6.8978 | -34.2151 | -12.1834
6.9039 | -34.2154 | -12.193
6.91 | -34.2157 | -12.203
6.9161 | -34.2161 | -12.2126
6.9222 | -34.2164 | -12.2225
6.9283 | -34.2168 | -12.2321
6.9344 | -34.2171 | -12.2421
6.9405 | -34.2174 | -12.2519
6.9466 | -34.2178 | -12.262
6.9528 | -34.2181 | -12.2721
6.9589 | -34.2185 | -12.2821
6.965 | -34.2188 | -12.2921
6.9711 | -34.2192 | -12.3021
6.9772 | -34.2195 | -12.312
6.9833 | -34.2198 | -12.3219
6.9894 | -34.2202 | -12.332
6.9955 | -34.2205 | -12.3417
7.0016 | -34.2208 | -12.3515
7.0077 | -34.2212 | -12.3614
7.0138 | -34.2215 | -12.3712
7.0199 | -34.2218 | -12.3809
7.026 | -34.2222 | -12.3907
7.0321 | -34.2225 | -12.4004
7.0382 | -34.2229 | -12.4099
7.0443 | -34.2232 | -12.42
7.0504 | -34.2235 | -12.4299
7.0565 | -34.2239 | -12.4398
7.0626 | -34.2242 | -12.4494
7.0687 | -34.2246 | -12.4588
7.0748 | -34.2249 | -12.4684
7.0809 | -34.2253 | -12.4782
7.087 | -34.2256 | -12.4881
7.0932 | -34.2259 | -12.498
7.0993 | -34.2262 | -12.5078
7.1054 | -34.2266 | -12.5175
7.1115 | -34.227 | -12.5271
7.1176 | -34.2273 | -12.5368
7.1237 | -34.2276 | -12.5466
7.1298 | -34.2279 | -12.5563
7.1359 | -34.2283 | -12.5661
7.142 | -34.2286 | -12.5756
7.1481 | -34.2289 | -12.5853
7.1542 | -34.2293 | -12.5951
7.1603 | -34.2296 | -12.6049
7.1664 | -34.23 | -12.6147
7.1725 | -34.2304 | -12.6246
7.1786 | -34.2307 | -12.6343
7.1847 | -34.231 | -12.6442
7.1908 | -34.2314 | -12.6541
7.1969 | -34.2317 | -12.6641
7.203 | -34.232 | -12.6737
7.2091 | -34.2324 | -12.6834
7.2152 | -34.2328 | -12.6932
7.2213 | -34.2331 | -12.7028
7.2274 | -34.2335 | -12.7126
7.2335 | -34.2338 | -12.7223
7.2397 | -34.2342 | -12.732
7.2458 | -34.2345 | -12.7416
7.2519 | -34.2349 | -12.7514
7.258 | -34.2352 | -12.7613
7.2641 | -34.2356 | -12.7711
7.2702 | -34.2359 | -12.781
7.2763 | -34.2363 | -12.7908
7.2824 | -34.2366 | -12.8008
7.2885 | -34.237 | -12.8106
7.2946 | -34.2373 | -12.8204
7.3007 | -34.2376 | -12.8302
7.3068 | -34.238 | -12.8403
7.3129 | -34.2383 | -12.8503
7.319 | -34.2387 | -12.86
7.3251 | -34.239 | -12.8696
7.3312 | -34.2394 | -12.8794
7.3373 | -34.2398 | -12.889
7.3434 | -34.2401 | -12.8986
7.3495 | -34.2405 | -12.9083
7.3556 | -34.2409 | -12.9182
7.3617 | -34.2412 | -12.9281
7.3678 | -34.2416 | -12.9378
7.3739 | -34.2419 | -12.9479
7.3801 | -34.2422 | -12.9579
7.3862 | -34.2426 | -12.9676
7.3923 | -34.243 | -12.9776
7.3984 | -34.2433 | -12.9871
7.4045 | -34.2437 | -12.9968
7.4106 | -34.2441 | -13.0066
7.4167 | -34.2444 | -13.0164
7.4228 | -34.2448 | -13.0263
7.4289 | -34.2452 | -13.0361
7.435 | -34.2455 | -13.0456
7.4411 | -34.2459 | -13.0555
7.4472 | -34.2462 | -13.0654
7.4533 | -34.2466 | -13.075
7.4594 | -34.247 | -13.085
7.4655 | -34.2473 | -13.0946
7.4716 | -34.2477 | -13.1046
7.4777 | -34.248 | -13.1143
7.4838 | -34.2484 | -13.124
7.4899 | -34.2487 | -13.1337
7.496 | -34.2491 | -13.1434
7.5021 | -34.2495 | -13.1531
7.5082 | -34.2498 | -13.1629
7.5143 | -34.2502 | -13.1727
7.5204 | -34.2505 | -13.1826
7.5266 | -34.2509 | -13.1923
7.5327 | -34.2512 | -13.202
7.5388 | -34.2516 | -13.2121
7.5449 | -34.252 | -13.2221
7.551 | -34.2523 | -13.2317
7.5571 | -34.2527 | -13.2416
7.5632 | -34.2531 | -13.2515
7.5693 | -34.2535 | -13.2612
7.5754 | -34.2538 | -13.2709
7.5815 | -34.2542 | -13.2807
7.5876 | -34.2545 | -13.2907
7.5937 | -34.2549 | -13.3007
7.5998 | -34.2553 | -13.3105
7.6059 | -34.2557 | -13.3205
7.612 | -34.2561 | -13.3302
7.6181 | -34.2564 | -13.34
7.6242 | -34.2568 | -13.3501
7.6303 | -34.2572 | -13.3602
7.6364 | -34.2575 | -13.3701
7.6425 | -34.2579 | -13.3799
7.6486 | -34.2582 | -13.3898
7.6547 | -34.2586 | -13.3995
7.6608 | -34.259 | -13.4093
7.667 | -34.2593 | -13.4193
7.6731 | -34.2597 | -13.4292
7.6792 | -34.2601 | -13.4392
7.6853 | -34.2604 | -13.4492
7.6914 | -34.2608 | -13.4591
7.6975 | -34.2612 | -13.4691
7.7036 | -34.2615 | -13.479
7.7097 | -34.2619 | -13.489
7.7158 | -34.2623 | -13.499
7.7219 | -34.2627 | -13.5087
7.728 | -34.263 | -13.5187
7.7341 | -34.2634 | -13.5287
7.7402 | -34.2638 | -13.5386
7.7463 | -34.2642 | -13.5486
7.7524 | -34.2645 | -13.5587
7.7585 | -34.2649 | -13.5686
7.7646 | -34.2652 | -13.5785
7.7707 | -34.2656 | -13.5886
7.7768 | -34.2659 | -13.5986
7.7829 | -34.2663 | -13.6088
7.789 | -34.2667 | -13.619
7.7951 | -34.267 | -13.6288
7.8012 | -34.2674 | -13.6387
7.8073 | -34.2678 | -13.6489
7.8135 | -34.2682 | -13.6588
7.8196 | -34.2686 | -13.669
7.8257 | -34.2689 | -13.6791
7.8318 | -34.2693 | -13.6891
7.8379 | -34.2697 | -13.699
7.844 | -34.27 | -13.7088
7.8501 | -34.2704 | -13.7189
7.8562 | -34.2708 | -13.729
7.8623 | -34.2712 | -13.7389
7.8684 | -34.2715 | -13.7489
7.8745 | -34.2719 | -13.7589
7.8806 | -34.2722 | -13.7691
7.8867 | -34.2726 | -13.7788
7.8928 | -34.273 | -13.7885
7.8989 | -34.2734 | -13.7985
7.905 | -34.2737 | -13.8085
7.9111 | -34.2741 | -13.8185
7.9172 | -34.2745 | -13.8285
7.9233 | -34.2748 | -13.8385
7.9294 | -34.2752 | -13.8486
7.9355 | -34.2756 | -13.8583
7.9416 | -34.2759 | -13.8683
7.9477 | -34.2763 | -13.8781
7.9539 | -34.2766 | -13.888
7.96 | -34.277 | -13.8981
7.9661 | -34.2774 | -13.9082
7.9722 | -34.2778 | -13.918
7.9783 | -34.2782 | -13.9282
7.9844 | -34.2786 | -13.9381
7.9905 | -34.2789 | -13.9484
7.9966 | -34.2793 | -13.9584
8.0027 | -34.2796 | -13.9684
8.0088 | -34.28 | -13.9784
8.0149 | -34.2803 | -13.9884
8.021 | -34.2807 | -13.9984
8.0271 | -34.2811 | -14.0087
8.0332 | -34.2814 | -14.0186
8.0393 | -34.2819 | -14.0288
8.0454 | -34.2822 | -14.039
8.0515 | -34.2826 | -14.049
8.0576 | -34.283 | -14.0593
8.0637 | -34.2833 | -14.0692
8.0698 | -34.2837 | -14.079
8.0759 | -34.284 | -14.0889
8.082 | -34.2844 | -14.0988
8.0881 | -34.2848 | -14.1089
8.0942 | -34.2852 | -14.1189
8.1004 | -34.2856 | -14.1291
8.1065 | -34.2859 | -14.1389
8.1126 | -34.2863 | -14.1489
8.1187 | -34.2866 | -14.159
8.1248 | -34.287 | -14.1688
8.1309 | -34.2874 | -14.1787
8.137 | -34.2877 | -14.1885
8.1431 | -34.2881 | -14.1988
8.1492 | -34.2885 | -14.2087
8.1553 | -34.2888 | -14.2188
8.1614 | -34.2892 | -14.2289
8.1675 | -34.2895 | -14.2389
8.1736 | -34.2899 | -14.2491
8.1797 | -34.2903 | -14.2592
8.1858 | -34.2906 | -14.2693
8.1919 | -34.291 | -14.2792
8.198 | -34.2914 | -14.289
8.2041 | -34.2918 | -14.2991
8.2102 | -34.2921 | -14.309
8.2163 | -34.2925 | -14.3189
8.2224 | -34.2929 | -14.3288
8.2285 | -34.2932 | -14.3385
8.2346 | -34.2936 | -14.3485
8.2408 | -34.294 | -14.3583
8.2469 | -34.2944 | -14.3684
8.253 | -34.2947 | -14.3782
8.2591 | -34.2951 | -14.3883
8.2652 | -34.2955 | -14.3983
8.2713 | -34.2959 | -14.4084
8.2774 | -34.2963 | -14.4185
8.2835 | -34.2966 | -14.4287
8.2896 | -34.297 | -14.4385
8.2957 | -34.2974 | -14.4486
8.3018 | -34.2978 | -14.4586
8.3079 | -34.2981 | -14.4684
8.314 | -34.2985 | -14.4784
8.3201 | -34.2989 | -14.4882
8.3262 | -34.2993 | -14.4982
8.3323 | -34.2996 | -14.5081
8.3384 | -34.3 | -14.5183
8.3445 | -34.3004 | -14.5281
8.3506 | -34.3008 | -14.5382
8.3567 | -34.3012 | -14.5483
8.3628 | -34.3015 | -14.5582
8.3689 | -34.3019 | -14.5681
8.375 | -34.3023 | -14.5782
8.3812 | -34.3026 | -14.588
8.3873 | -34.303 | -14.598
8.3934 | -34.3034 | -14.6081
8.3995 | -34.3037 | -14.6179
8.4056 | -34.3041 | -14.6278
8.4117 | -34.3044 | -14.6379
8.4178 | -34.3048 | -14.6479
8.4239 | -34.3052 | -14.6577
8.43 | -34.3055 | -14.6676
8.4361 | -34.3059 | -14.6777
8.4422 | -34.3063 | -14.6875
8.4483 | -34.3066 | -14.6973
8.4544 | -34.307 | -14.7073
8.4605 | -34.3074 | -14.7171
8.4666 | -34.3077 | -14.7273
8.4727 | -34.3081 | -14.7373
8.4788 | -34.3085 | -14.7473
8.4849 | -34.3088 | -14.7572
8.491 | -34.3092 | -14.7672
8.4971 | -34.3096 | -14.7768
8.5032 | -34.3099 | -14.7867
8.5093 | -34.3103 | -14.7964
8.5154 | -34.3107 | -14.8065
8.5215 | -34.3111 | -14.8163
8.5277 | -34.3114 | -14.8262
8.5338 | -34.3118 | -14.836
8.5399 | -34.3121 | -14.8459
8.546 | -34.3125 | -14.8556
8.5521 | -34.3129 | -14.8656
8.5582 | -34.3133 | -14.8755
8.5643 | -34.3136 | -14.8851
8.5704 | -34.314 | -14.8947
8.5765 | -34.3144 | -14.9048
8.5826 | -34.3148 | -14.9143
8.5887 | -34.3152 | -14.9242
8.5948 | -34.3156 | -14.934
8.6009 | -34.3159 | -14.9437
8.607 | -34.3163 | -14.9537
8.6131 | -34.3167 | -14.9637
8.6192 | -34.3171 | -14.9733
8.6253 | -34.3174 | -14.9831
8.6314 | -34.3178 | -14.9928
8.6375 | -34.3181 | -15.0028
8.6436 | -34.3185 | -15.0126
8.6497 | -34.3189 | -15.0223
8.6558 | -34.3192 | -15.0321
8.6619 | -34.3196 | -15.042
8.6681 | -34.3199 | -15.052
8.6742 | -34.3203 | -15.0617
8.6803 | -34.3207 | -15.0716
8.6864 | -34.3211 | -15.0811
8.6925 | -34.3214 | -15.0907
8.6986 | -34.3218 | -15.1006
8.7047 | -34.3221 | -15.1107
8.7108 | -34.3225 | -15.1205
8.7169 | -34.3229 | -15.1303
8.723 | -34.3232 | -15.1399
8.7291 | -34.3236 | -15.1495
8.7352 | -34.324 | -15.1595
8.7413 | -34.3243 | -15.1694
8.7474 | -34.3247 | -15.1789
8.7535 | -34.3251 | -15.1887
8.7596 | -34.3255 | -15.1985
8.7657 | -34.3259 | -15.2084
8.7718 | -34.3262 | -15.2183
8.7779 | -34.3266 | -15.2282
8.784 | -34.327 | -15.2379
8.7901 | -34.3274 | -15.2477
8.7962 | -34.3277 | -15.2572
8.8023 | -34.3281 | -15.2669
8.8084 | -34.3285 | -15.2766
8.8146 | -34.3288 | -15.2863
8.8207 | -34.3292 | -15.296
8.8268 | -34.3296 | -15.3054
8.8329 | -34.33 | -15.3152
8.839 | -34.3304 | -15.3249
8.8451 | -34.3307 | -15.3345
8.8512 | -34.3311 | -15.3442
8.8573 | -34.3314 | -15.3539
8.8634 | -34.3318 | -15.3634
8.8695 | -34.3322 | -15.3733
8.8756 | -34.3325 | -15.383
8.8817 | -34.3329 | -15.3926
8.8878 | -34.3333 | -15.4026
8.8939 | -34.3337 | -15.4123
8.9 | -34.3341 | -15.422
8.9061 | -34.3344 | -15.4318
8.9122 | -34.3348 | -15.4415
8.9183 | -34.3352 | -15.4512
8.9244 | -34.3356 | -15.4611
8.9305 | -34.336 | -15.4706
8.9366 | -34.3364 | -15.4803
8.9427 | -34.3368 | -15.4901
8.9488 | -34.3372 | -15.4997
8.955 | -34.3375 | -15.5094
8.9611 | -34.3379 | -15.5193
8.9672 | -34.3384 | -15.5288
8.9733 | -34.3388 | -15.5384
8.9794 | -34.3392 | -15.548
8.9855 | -34.3395 | -15.5578
8.9916 | -34.3399 | -15.5676
8.9977 | -34.3403 | -15.5772
9.0038 | -34.3407 | -15.5867
9.0099 | -34.3411 | -15.5964
9.016 | -34.3415 | -15.6061
9.0221 | -34.3419 | -15.6156
9.0282 | -34.3423 | -15.6251
9.0343 | -34.3427 | -15.6348
9.0404 | -34.3431 | -15.6445
9.0465 | -34.3435 | -15.6543
9.0526 | -34.3439 | -15.6639
9.0587 | -34.3443 | -15.6736
9.0648 | -34.3447 | -15.6831
9.0709 | -34.3451 | -15.6927
9.077 | -34.3455 | -15.7025
9.0831 | -34.3459 | -15.7123
9.0892 | -34.3463 | -15.722
9.0953 | -34.3467 | -15.7318
9.1015 | -34.3471 | -15.7415
9.1076 | -34.3475 | -15.7509
9.1137 | -34.3479 | -15.7605
9.1198 | -34.3483 | -15.7702
9.1259 | -34.3487 | -15.7799
9.132 | -34.3491 | -15.7897
9.1381 | -34.3495 | -15.7991
9.1442 | -34.35 | -15.8084
9.1503 | -34.3504 | -15.8182
9.1564 | -34.3508 | -15.8278
9.1625 | -34.3512 | -15.8374
9.1686 | -34.3516 | -15.847
9.1747 | -34.352 | -15.8566
9.1808 | -34.3524 | -15.8661
9.1869 | -34.3528 | -15.8757
9.193 | -34.3532 | -15.8853
9.1991 | -34.3536 | -15.895
9.2052 | -34.354 | -15.9046
9.2113 | -34.3544 | -15.9143
9.2174 | -34.3549 | -15.9239
9.2235 | -34.3553 | -15.9335
9.2296 | -34.3557 | -15.943
9.2357 | -34.3561 | -15.9528
9.2419 | -34.3565 | -15.9623
9.248 | -34.357 | -15.9719
9.2541 | -34.3574 | -15.9815
9.2602 | -34.3578 | -15.9913
9.2663 | -34.3582 | -16.0007
9.2724 | -34.3587 | -16.0102
9.2785 | -34.3591 | -16.0198
9.2846 | -34.3595 | -16.0292
9.2907 | -34.3599 | -16.0388
9.2968 | -34.3603 | -16.0485
9.3029 | -34.3607 | -16.0583
9.309 | -34.3611 | -16.0679
9.3151 | -34.3615 | -16.0775
9.3212 | -34.362 | -16.0869
9.3273 | -34.3624 | -16.0964
9.3334 | -34.3628 | -16.1059
9.3395 | -34.3632 | -16.1153
9.3456 | -34.3636 | -16.1247
9.3517 | -34.364 | -16.1344
9.3578 | -34.3644 | -16.144
9.3639 | -34.3649 | -16.1534
9.37 | -34.3653 | -16.163
9.3761 | -34.3657 | -16.1728
9.3822 | -34.3662 | -16.1824
9.3884 | -34.3666 | -16.1917
9.3945 | -34.367 | -16.2013
9.4006 | -34.3675 | -16.211
9.4067 | -34.3679 | -16.2206
9.4128 | -34.3683 | -16.2303
9.4189 | -34.3687 | -16.2401
9.425 | -34.3692 | -16.2494
9.4311 | -34.3696 | -16.2591
9.4372 | -34.37 | -16.2686
9.4433 | -34.3705 | -16.2782
9.4494 | -34.3709 | -16.2875
9.4555 | -34.3713 | -16.2972
9.4616 | -34.3718 | -16.3068
9.4677 | -34.3722 | -16.3163
9.4738 | -34.3726 | -16.3259
9.4799 | -34.3731 | -16.3355
9.486 | -34.3735 | -16.345
9.4921 | -34.3739 | -16.3547
9.4982 | -34.3743 | -16.3639
9.5043 | -34.3748 | -16.3735
9.5104 | -34.3752 | -16.3828
9.5165 | -34.3756 | -16.3923
9.5226 | -34.376 | -16.402
9.5288 | -34.3765 | -16.4113
9.5349 | -34.3769 | -16.4209
9.541 | -34.3773 | -16.4303
9.5471 | -34.3778 | -16.4397
9.5532 | -34.3782 | -16.4494
9.5593 | -34.3786 | -16.459
9.5654 | -34.3791 | -16.4685
9.5715 | -34.3795 | -16.478
9.5776 | -34.3799 | -16.4877
9.5837 | -34.3803 | -16.4972
9.5898 | -34.3807 | -16.5069
9.5959 | -34.3812 | -16.5164
9.602 | -34.3816 | -16.5258
9.6081 | -34.382 | -16.5353
9.6142 | -34.3825 | -16.5448
9.6203 | -34.3829 | -16.5542
9.6264 | -34.3833 | -16.5637
9.6325 | -34.3837 | -16.5733
9.6386 | -34.3842 | -16.5826
9.6447 | -34.3846 | -16.5922
9.6508 | -34.385 | -16.6016
9.6569 | -34.3854 | -16.6111
9.663 | -34.3859 | -16.6206
9.6691 | -34.3863 | -16.6302
9.6753 | -34.3868 | -16.6397
9.6814 | -34.3872 | -16.649
9.6875 | -34.3876 | -16.6583
9.6936 | -34.3881 | -16.6676
9.6997 | -34.3885 | -16.6773
9.7058 | -34.389 | -16.6867
9.7119 | -34.3894 | -16.6963
9.718 | -34.3899 | -16.7059
9.7241 | -34.3903 | -16.7154
9.7302 | -34.3907 | -16.725
9.7363 | -34.3912 | -16.7344
9.7424 | -34.3916 | -16.7438
9.7485 | -34.392 | -16.7534
9.7546 | -34.3924 | -16.7626
9.7607 | -34.3929 | -16.7723
9.7668 | -34.3933 | -16.7819
9.7729 | -34.3938 | -16.7914
9.779 | -34.3942 | -16.8006
9.7851 | -34.3947 | -16.8103
9.7912 | -34.3951 | -16.8197
9.7973 | -34.3956 | -16.8293
9.8034 | -34.396 | -16.8386
9.8095 | -34.3965 | -16.8484
9.8157 | -34.3969 | -16.8579
9.8218 | -34.3974 | -16.8675
9.8279 | -34.3978 | -16.8769
9.834 | -34.3982 | -16.8865
9.8401 | -34.3986 | -16.8959
9.8462 | -34.3991 | -16.9051
9.8523 | -34.3995 | -16.9146
9.8584 | -34.3999 | -16.9241
9.8645 | -34.4004 | -16.9337
9.8706 | -34.4008 | -16.9432
9.8767 | -34.4012 | -16.9526
9.8828 | -34.4016 | -16.962
9.8889 | -34.4021 | -16.9713
9.895 | -34.4025 | -16.9808
9.9011 | -34.4029 | -16.9903
9.9072 | -34.4034 | -16.9997
9.9133 | -34.4038 | -17.0092
9.9194 | -34.4042 | -17.0186
9.9255 | -34.4046 | -17.028
9.9316 | -34.4051 | -17.0374
9.9377 | -34.4055 | -17.0468
9.9438 | -34.4059 | -17.0563
9.9499 | -34.4064 | -17.066
9.956 | -34.4068 | -17.0753
9.9622 | -34.4073 | -17.0848
9.9683 | -34.4077 | -17.0943
9.9744 | -34.4081 | -17.1039
9.9805 | -34.4086 | -17.1134
9.9866 | -34.409 | -17.1229
9.9927 | -34.4094 | -17.1322
9.9988 | -34.4098 | -17.1418
10.0049 | -34.4103 | -17.1512
10.011 | -34.4107 | -17.1608
10.0171 | -34.4112 | -17.1702
10.0232 | -34.4116 | -17.1796
10.0293 | -34.4121 | -17.1892
10.0354 | -34.4125 | -17.1986
10.0415 | -34.413 | -17.2081
10.0476 | -34.4134 | -17.2177
10.0537 | -34.4138 | -17.2272
10.0598 | -34.4143 | -17.2366
10.0659 | -34.4147 | -17.2462
10.072 | -34.4151 | -17.2555
10.0781 | -34.4155 | -17.2649
10.0842 | -34.416 | -17.2744
10.0903 | -34.4164 | -17.2838
10.0964 | -34.4169 | -17.293
10.1026 | -34.4173 | -17.3025
10.1087 | -34.4177 | -17.3119
10.1148 | -34.4181 | -17.3214
10.1209 | -34.4185 | -17.3307
10.127 | -34.419 | -17.3401
10.1331 | -34.4194 | -17.3495
10.1392 | -34.4198 | -17.3587
10.1453 | -34.4203 | -17.3681
10.1514 | -34.4207 | -17.3777
10.1575 | -34.4211 | -17.3872
10.1636 | -34.4216 | -17.3964
10.1697 | -34.422 | -17.4059
10.1758 | -34.4225 | -17.4151
10.1819 | -34.4229 | -17.4245
10.188 | -34.4233 | -17.434
10.1941 | -34.4237 | -17.4435
10.2002 | -34.4241 | -17.4531
10.2063 | -34.4246 | -17.4624
10.2124 | -34.425 | -17.4719
10.2185 | -34.4254 | -17.4815
10.2246 | -34.4259 | -17.4907
10.2307 | -34.4263 | -17.5002
10.2368 | -34.4267 | -17.5094
10.2429 | -34.4271 | -17.5187
10.2491 | -34.4276 | -17.5282
10.2552 | -34.428 | -17.5377
10.2613 | -34.4284 | -17.5472
10.2674 | -34.4288 | -17.5567
10.2735 | -34.4293 | -17.5663
10.2796 | -34.4297 | -17.5756
10.2857 | -34.4301 | -17.5851
10.2918 | -34.4306 | -17.5947
10.2979 | -34.431 | -17.6041
10.304 | -34.4314 | -17.6134
10.3101 | -34.4319 | -17.6228
10.3162 | -34.4323 | -17.6324
10.3223 | -34.4327 | -17.6416
10.3284 | -34.4331 | -17.6509
10.3345 | -34.4336 | -17.6603
10.3406 | -34.434 | -17.6696
10.3467 | -34.4344 | -17.679
10.3528 | -34.4348 | -17.6882
10.3589 | -34.4352 | -17.6977
10.365 | -34.4357 | -17.7072
10.3711 | -34.4361 | -17.7167
10.3772 | -34.4365 | -17.726
10.3833 | -34.437 | -17.7355
10.3895 | -34.4374 | -17.7448
10.3956 | -34.4379 | -17.7542
10.4017 | -34.4383 | -17.7635
10.4078 | -34.4387 | -17.7729
10.4139 | -34.4391 | -17.7824
10.42 | -34.4396 | -17.792
10.4261 | -34.44 | -17.8014
10.4322 | -34.4404 | -17.8108
10.4383 | -34.4408 | -17.8202
10.4444 | -34.4412 | -17.8294
10.4505 | -34.4416 | -17.8389
10.4566 | -34.4421 | -17.8481
10.4627 | -34.4425 | -17.8575
10.4688 | -34.4429 | -17.8667
10.4749 | -34.4433 | -17.8762
10.481 | -34.4438 | -17.8856
10.4871 | -34.4442 | -17.8949
10.4932 | -34.4446 | -17.9042
10.4993 | -34.4451 | -17.9137
10.5054 | -34.4455 | -17.9231
10.5115 | -34.4459 | -17.9324
10.5176 | -34.4464 | -17.9417
10.5237 | -34.4468 | -17.9513
10.5298 | -34.4472 | -17.9604
10.536 | -34.4476 | -17.9699
10.5421 | -34.4481 | -17.9792
10.5482 | -34.4485 | -17.9886
10.5543 | -34.4489 | -17.9979
10.5604 | -34.4493 | -18.007
10.5665 | -34.4498 | -18.0162
10.5726 | -34.4502 | -18.0256
10.5787 | -34.4506 | -18.0349
10.5848 | -34.451 | -18.0439
10.5909 | -34.4514 | -18.0534
10.597 | -34.4519 | -18.0628
10.6031 | -34.4523 | -18.0724
10.6092 | -34.4527 | -18.0817
10.6153 | -34.4531 | -18.0912
10.6214 | -34.4536 | -18.1005
10.6275 | -34.454 | -18.1098
10.6336 | -34.4544 | -18.1192
10.6397 | -34.4548 | -18.1285
10.6458 | -34.4552 | -18.1378
10.6519 | -34.4556 | -18.1473
10.658 | -34.4561 | -18.1566
10.6641 | -34.4565 | -18.1661
10.6702 | -34.4569 | -18.1754
10.6764 | -34.4573 | -18.1849
10.6825 | -34.4578 | -18.1941
10.6886 | -34.4582 | -18.2033
10.6947 | -34.4586 | -18.2127
10.7008 | -34.459 | -18.2221
10.7069 | -34.4595 | -18.2314
10.713 | -34.4599 | -18.2407
10.7191 | -34.4603 | -18.2501
10.7252 | -34.4607 | -18.2591
10.7313 | -34.4611 | -18.2684
10.7374 | -34.4615 | -18.2777
10.7435 | -34.4619 | -18.2872
10.7496 | -34.4623 | -18.2963
10.7557 | -34.4627 | -18.3058
10.7618 | -34.4632 | -18.3149
10.7679 | -34.4636 | -18.3242
10.774 | -34.464 | -18.3336
10.7801 | -34.4644 | -18.343
10.7862 | -34.4648 | -18.3521
10.7923 | -34.4652 | -18.3617
10.7984 | -34.4657 | -18.3707
10.8045 | -34.4661 | -18.3798
10.8106 | -34.4665 | -18.3893
10.8168 | -34.4669 | -18.3986
10.8229 | -34.4674 | -18.4079
10.829 | -34.4678 | -18.4171
10.8351 | -34.4682 | -18.4262
10.8412 | -34.4686 | -18.4354
10.8473 | -34.4691 | -18.4445
10.8534 | -34.4695 | -18.4537
10.8595 | -34.4699 | -18.4631
10.8656 | -34.4703 | -18.4723
10.8717 | -34.4707 | -18.4816
10.8778 | -34.4712 | -18.4908
10.8839 | -34.4716 | -18.5
10.89 | -34.472 | -18.5094
10.8961 | -34.4724 | -18.5186
10.9022 | -34.4728 | -18.5277
10.9083 | -34.4733 | -18.5368
10.9144 | -34.4737 | -18.5461
10.9205 | -34.4741 | -18.5553
10.9266 | -34.4745 | -18.5645
10.9327 | -34.4749 | -18.5738
10.9388 | -34.4754 | -18.5829
10.9449 | -34.4758 | -18.592
10.951 | -34.4762 | -18.6013
10.9571 | -34.4766 | -18.6105
10.9633 | -34.477 | -18.6199
10.9694 | -34.4774 | -18.6294
10.9755 | -34.4779 | -18.6387
10.9816 | -34.4783 | -18.6479
10.9877 | -34.4787 | -18.657
10.9938 | -34.4791 | -18.6665
10.9999 | -34.4796 | -18.6755
11.006 | -34.48 | -18.6845
11.0121 | -34.4804 | -18.6937
11.0182 | -34.4809 | -18.7029
11.0243 | -34.4813 | -18.7121
11.0304 | -34.4817 | -18.7214
11.0365 | -34.4822 | -18.7305
11.0426 | -34.4826 | -18.7397
11.0487 | -34.483 | -18.749
11.0548 | -34.4835 | -18.7581
11.0609 | -34.4839 | -18.7674
11.067 | -34.4843 | -18.7767
11.0731 | -34.4848 | -18.7858
11.0792 | -34.4852 | -18.7949
11.0853 | -34.4856 | -18.8041
11.0914 | -34.4861 | -18.8134
11.0975 | -34.4865 | -18.8227
11.1037 | -34.4869 | -18.8317
11.1098 | -34.4873 | -18.8411
11.1159 | -34.4877 | -18.8501
11.122 | -34.4881 | -18.8592
11.1281 | -34.4886 | -18.8684
11.1342 | -34.489 | -18.8777
11.1403 | -34.4894 | -18.887
11.1464 | -34.4899 | -18.8962
11.1525 | -34.4903 | -18.9053
11.1586 | -34.4907 | -18.9146
11.1647 | -34.4912 | -18.9238
11.1708 | -34.4916 | -18.9329
11.1769 | -34.492 | -18.9421
11.183 | -34.4924 | -18.951
11.1891 | -34.4929 | -18.9602
11.1952 | -34.4933 | -18.9693
11.2013 | -34.4938 | -18.9786
11.2074 | -34.4942 | -18.9877
11.2135 | -34.4946 | -18.9971
11.2196 | -34.495 | -19.0063
11.2257 | -34.4955 | -19.0155
11.2318 | -34.4959 | -19.0246
11.2379 | -34.4964 | -19.0339
11.244 | -34.4968 | -19.0432
11.2502 | -34.4973 | -19.0523
11.2563 | -34.4977 | -19.0615
11.2624 | -34.4982 | -19.0707
11.2685 | -34.4986 | -19.0797
11.2746 | -34.499 | -19.0888
11.2807 | -34.4995 | -19.0981
11.2868 | -34.4999 | -19.1074
11.2929 | -34.5004 | -19.1162
11.299 | -34.5008 | -19.1254
11.3051 | -34.5012 | -19.1347
11.3112 | -34.5017 | -19.144
11.3173 | -34.5021 | -19.1529
11.3234 | -34.5026 | -19.1619
11.3295 | -34.503 | -19.1711
11.3356 | -34.5034 | -19.1803
11.3417 | -34.5039 | -19.1893
11.3478 | -34.5043 | -19.1987
11.3539 | -34.5047 | -19.2079
11.36 | -34.5051 | -19.217
11.3661 | -34.5056 | -19.226
11.3722 | -34.506 | -19.2349
11.3783 | -34.5065 | -19.2441
11.3844 | -34.5069 | -19.2533
11.3906 | -34.5074 | -19.2624
11.3967 | -34.5078 | -19.2715
11.4028 | -34.5083 | -19.2808
11.4089 | -34.5087 | -19.2898
11.415 | -34.5091 | -19.2988
11.4211 | -34.5095 | -19.308
11.4272 | -34.51 | -19.3171
11.4333 | -34.5104 | -19.3259
11.4394 | -34.5108 | -19.335
11.4455 | -34.5112 | -19.3439
11.4516 | -34.5117 | -19.3532
11.4577 | -34.5121 | -19.3623
11.4638 | -34.5126 | -19.3712
11.4699 | -34.513 | -19.3804
11.476 | -34.5134 | -19.3896
11.4821 | -34.5139 | -19.3987
11.4882 | -34.5144 | -19.4079
11.4943 | -34.5148 | -19.4172
11.5004 | -34.5153 | -19.4261
11.5065 | -34.5157 | -19.4351
11.5126 | -34.5162 | -19.4441
11.5187 | -34.5166 | -19.4532
11.5248 | -34.5171 | -19.4626
11.5309 | -34.5175 | -19.4717
11.5371 | -34.518 | -19.4805
11.5432 | -34.5184 | -19.4898
11.5493 | -34.5188 | -19.499
11.5554 | -34.5193 | -19.5081
11.5615 | -34.5197 | -19.5174
11.5676 | -34.5202 | -19.5265
11.5737 | -34.5207 | -19.5355
11.5798 | -34.5211 | -19.5444
11.5859 | -34.5216 | -19.5535
11.592 | -34.522 | -19.5625
11.5981 | -34.5224 | -19.5716
11.6042 | -34.5229 | -19.5806
11.6103 | -34.5233 | -19.5896
11.6164 | -34.5238 | -19.5985
11.6225 | -34.5242 | -19.6075
11.6286 | -34.5247 | -19.6166
11.6347 | -34.5251 | -19.6256
11.6408 | -34.5255 | -19.6345
11.6469 | -34.526 | -19.6435
11.653 | -34.5264 | -19.6524
11.6591 | -34.5269 | -19.6616
11.6652 | -34.5273 | -19.6705
11.6713 | -34.5278 | -19.6795
11.6775 | -34.5282 | -19.6884
11.6836 | -34.5287 | -19.6974
11.6897 | -34.5291 | -19.7065
11.6958 | -34.5296 | -19.7156
11.7019 | -34.53 | -19.7249
11.708 | -34.5305 | -19.7339
11.7141 | -34.531 | -19.7427
11.7202 | -34.5314 | -19.7517
11.7263 | -34.5318 | -19.7608
11.7324 | -34.5323 | -19.77
11.7385 | -34.5327 | -19.7791
11.7446 | -34.5332 | -19.788
11.7507 | -34.5337 | -19.7969
11.7568 | -34.5341 | -19.8059
11.7629 | -34.5345 | -19.8151
11.769 | -34.535 | -19.8239
11.7751 | -34.5354 | -19.833
11.7812 | -34.5359 | -19.8419
11.7873 | -34.5364 | -19.8509
11.7934 | -34.5368 | -19.8599
11.7995 | -34.5373 | -19.8689
11.8056 | -34.5377 | -19.8777
11.8117 | -34.5381 | -19.8868
11.8178 | -34.5386 | -19.8957
11.824 | -34.5391 | -19.9046
11.8301 | -34.5396 | -19.9136
11.8362 | -34.54 | -19.9226
11.8423 | -34.5404 | -19.9315
11.8484 | -34.5409 | -19.9406
11.8545 | -34.5414 | -19.9494
11.8606 | -34.5418 | -19.9583
11.8667 | -34.5423 | -19.9674
11.8728 | -34.5427 | -19.9762
11.8789 | -34.5432 | -19.9852
11.885 | -34.5436 | -19.9942
11.8911 | -34.5441 | -20.0029
11.8972 | -34.5446 | -20.0118
11.9033 | -34.545 | -20.0207
11.9094 | -34.5455 | -20.0296
11.9155 | -34.5459 | -20.0386
11.9216 | -34.5464 | -20.0474
11.9277 | -34.5469 | -20.0563
11.9338 | -34.5473 | -20.0653
11.9399 | -34.5478 | -20.0743
11.946 | -34.5482 | -20.0831
11.9521 | -34.5487 | -20.0919
11.9582 | -34.5492 | -20.1007
11.9644 | -34.5496 | -20.1096
11.9705 | -34.5501 | -20.1185
11.9766 | -34.5506 | -20.1274
11.9827 | -34.5511 | -20.1363
11.9888 | -34.5516 | -20.1451
11.9949 | -34.552 | -20.1538
12.001 | -34.5525 | -20.1628
12.0071 | -34.5529 | -20.1715
12.0132 | -34.5534 | -20.1803
12.0193 | -34.5538 | -20.1894
12.0254 | -34.5543 | -20.1983
12.0315 | -34.5548 | -20.207
12.0376 | -34.5552 | -20.2158
12.0437 | -34.5557 | -20.2247
12.0498 | -34.5561 | -20.2335
12.0559 | -34.5566 | -20.2423
12.062 | -34.5571 | -20.2509
12.0681 | -34.5576 | -20.2599
12.0742 | -34.558 | -20.2687
12.0803 | -34.5585 | -20.2778
12.0864 | -34.559 | -20.2867
12.0925 | -34.5594 | -20.2955
12.0986 | -34.5599 | -20.3043
12.1047 | -34.5604 | -20.3133
12.1109 | -34.5609 | -20.322
12.117 | -34.5613 | -20.3308
12.1231 | -34.5618 | -20.3396
12.1292 | -34.5622 | -20.3485
12.1353 | -34.5627 | -20.3572
12.1414 | -34.5632 | -20.3661
12.1475 | -34.5637 | -20.3751
12.1536 | -34.5641 | -20.384
12.1597 | -34.5646 | -20.3928
12.1658 | -34.5651 | -20.4016
12.1719 | -34.5655 | -20.4102
12.178 | -34.566 | -20.4191
12.1841 | -34.5665 | -20.428
12.1902 | -34.5669 | -20.4369
12.1963 | -34.5674 | -20.4458
12.2024 | -34.5679 | -20.4545
12.2085 | -34.5684 | -20.4633
12.2146 | -34.5688 | -20.4722
12.2207 | -34.5693 | -20.481
12.2268 | -34.5698 | -20.4897
12.2329 | -34.5702 | -20.4985
12.239 | -34.5708 | -20.507
12.2451 | -34.5712 | -20.5158
12.2513 | -34.5717 | -20.5248
12.2574 | -34.5721 | -20.5336
12.2635 | -34.5726 | -20.5424
12.2696 | -34.5731 | -20.5511
12.2757 | -34.5736 | -20.5597
12.2818 | -34.5741 | -20.5686
12.2879 | -34.5745 | -20.5775
12.294 | -34.575 | -20.5862
12.3001 | -34.5755 | -20.5951
12.3062 | -34.576 | -20.604
12.3123 | -34.5765 | -20.6126
12.3184 | -34.577 | -20.6215
12.3245 | -34.5775 | -20.6302
12.3306 | -34.5779 | -20.6389
12.3367 | -34.5784 | -20.6476
12.3428 | -34.5789 | -20.6563
12.3489 | -34.5794 | -20.6648
12.355 | -34.5798 | -20.6738
12.3611 | -34.5803 | -20.6826
12.3672 | -34.5808 | -20.6913
12.3733 | -34.5812 | -20.7
12.3794 | -34.5817 | -20.7084
12.3855 | -34.5822 | -20.7171
12.3916 | -34.5827 | -20.7259
12.3978 | -34.5832 | -20.7348
12.4039 | -34.5836 | -20.7436
12.41 | -34.5841 | -20.7524
12.4161 | -34.5846 | -20.7608
12.4222 | -34.5851 | -20.7694
12.4283 | -34.5855 | -20.7781
12.4344 | -34.586 | -20.7869
12.4405 | -34.5865 | -20.7957
12.4466 | -34.587 | -20.8043
12.4527 | -34.5874 | -20.813
12.4588 | -34.5879 | -20.8217
12.4649 | -34.5884 | -20.8305
12.471 | -34.5889 | -20.8393
12.4771 | -34.5893 | -20.8479
12.4832 | -34.5898 | -20.8566
12.4893 | -34.5903 | -20.8652
12.4954 | -34.5908 | -20.8738
12.5015 | -34.5913 | -20.8825
12.5076 | -34.5918 | -20.8914
12.5137 | -34.5923 | -20.9001
12.5198 | -34.5927 | -20.9087
12.5259 | -34.5932 | -20.9176
12.532 | -34.5936 | -20.9262
12.5382 | -34.5941 | -20.9349
12.5443 | -34.5946 | -20.9435
12.5504 | -34.595 | -20.9523
12.5565 | -34.5956 | -20.9608
12.5626 | -34.5961 | -20.9696
12.5687 | -34.5965 | -20.9784
12.5748 | -34.597 | -20.9872
12.5809 | -34.5974 | -20.9958
12.587 | -34.5979 | -21.0047
12.5931 | -34.5985 | -21.0132
12.5992 | -34.599 | -21.0218
12.6053 | -34.5994 | -21.0305
12.6114 | -34.5999 | -21.0391
12.6175 | -34.6004 | -21.0479
12.6236 | -34.6008 | -21.0566
12.6297 | -34.6014 | -21.0652
12.6358 | -34.6018 | -21.0737
12.6419 | -34.6023 | -21.0825
12.648 | -34.6028 | -21.0913
12.6541 | -34.6033 | -21.0999
12.6602 | -34.6038 | -21.1082
12.6663 | -34.6042 | -21.1168
12.6724 | -34.6047 | -21.1254
12.6785 | -34.6052 | -21.134
12.6847 | -34.6057 | -21.1426
12.6908 | -34.6062 | -21.1513
12.6969 | -34.6067 | -21.1598
12.703 | -34.6072 | -21.1685
12.7091 | -34.6076 | -21.1772
12.7152 | -34.6081 | -21.1857
12.7213 | -34.6086 | -21.1944
12.7274 | -34.609 | -21.203
12.7335 | -34.6095 | -21.2115
12.7396 | -34.61 | -21.2199
12.7457 | -34.6105 | -21.2286
12.7518 | -34.6109 | -21.2372
12.7579 | -34.6114 | -21.2459
12.764 | -34.6119 | -21.2546
12.7701 | -34.6123 | -21.2631
12.7762 | -34.6128 | -21.2716
12.7823 | -34.6133 | -21.2803
12.7884 | -34.6138 | -21.2888
12.7945 | -34.6143 | -21.2976
12.8006 | -34.6148 | -21.3063
12.8067 | -34.6153 | -21.3149
12.8128 | -34.6158 | -21.3233
12.8189 | -34.6162 | -21.3319
12.8251 | -34.6167 | -21.3406
12.8312 | -34.6172 | -21.3493
12.8373 | -34.6177 | -21.3579
12.8434 | -34.6181 | -21.3666
12.8495 | -34.6186 | -21.3751
12.8556 | -34.6191 | -21.3836
12.8617 | -34.6196 | -21.3922
12.8678 | -34.62 | -21.4009
12.8739 | -34.6205 | -21.4095
12.88 | -34.621 | -21.418
12.8861 | -34.6215 | -21.4265
12.8922 | -34.622 | -21.4352
12.8983 | -34.6224 | -21.4437
12.9044 | -34.6229 | -21.452
12.9105 | -34.6234 | -21.4604
12.9166 | -34.6239 | -21.469
12.9227 | -34.6243 | -21.4777
12.9288 | -34.6248 | -21.4862
12.9349 | -34.6253 | -21.4948
12.941 | -34.6258 | -21.5033
12.9471 | -34.6263 | -21.5119
12.9532 | -34.6268 | -21.5206
12.9593 | -34.6273 | -21.5294
12.9654 | -34.6278 | -21.538
12.9716 | -34.6282 | -21.5468
12.9777 | -34.6287 | -21.5554
12.9838 | -34.6292 | -21.564
12.9899 | -34.6297 | -21.5727
12.996 | -34.6302 | -21.5812
13.0021 | -34.6307 | -21.5897
13.0082 | -34.6312 | -21.5982
13.0143 | -34.6317 | -21.6064
13.0204 | -34.6322 | -21.6148
13.0265 | -34.6326 | -21.6232
13.0326 | -34.6331 | -21.6318
13.0387 | -34.6336 | -21.6405
13.0448 | -34.6341 | -21.6491
13.0509 | -34.6346 | -21.6576
13.057 | -34.6351 | -21.6662
13.0631 | -34.6356 | -21.6747
13.0692 | -34.636 | -21.6831
13.0753 | -34.6365 | -21.6919
13.0814 | -34.637 | -21.7006
13.0875 | -34.6375 | -21.7091
13.0936 | -34.638 | -21.7177
13.0997 | -34.6384 | -21.7261
13.1058 | -34.6389 | -21.7346
13.112 | -34.6394 | -21.7432
13.1181 | -34.6399 | -21.7517
13.1242 | -34.6404 | -21.7604
13.1303 | -34.6409 | -21.769
13.1364 | -34.6414 | -21.7775
13.1425 | -34.6418 | -21.786
13.1486 | -34.6423 | -21.7944
13.1547 | -34.6428 | -21.803
13.1608 | -34.6433 | -21.8114
13.1669 | -34.6438 | -21.8198
13.173 | -34.6443 | -21.8284
13.1791 | -34.6447 | -21.8369
13.1852 | -34.6452 | -21.8453
13.1913 | -34.6457 | -21.8537
13.1974 | -34.6461 | -21.8623
13.2035 | -34.6466 | -21.8708
13.2096 | -34.6471 | -21.8793
13.2157 | -34.6476 | -21.8877
13.2218 | -34.6481 | -21.8961
13.2279 | -34.6486 | -21.9044
13.234 | -34.6491 | -21.9128
13.2401 | -34.6496 | -21.9212
13.2462 | -34.65 | -21.9296
13.2524 | -34.6505 | -21.9383
13.2585 | -34.651 | -21.9466
13.2646 | -34.6515 | -21.9551
13.2707 | -34.652 | -21.9636
13.2768 | -34.6525 | -21.9723
13.2829 | -34.6529 | -21.9808
13.289 | -34.6534 | -21.9891
13.2951 | -34.6539 | -21.9975
13.3012 | -34.6544 | -22.006
13.3073 | -34.6549 | -22.0146
13.3134 | -34.6553 | -22.023
13.3195 | -34.6558 | -22.0314
13.3256 | -34.6563 | -22.0398
13.3317 | -34.6568 | -22.0483
13.3378 | -34.6572 | -22.0569
13.3439 | -34.6577 | -22.0655
13.35 | -34.6582 | -22.0739
13.3561 | -34.6587 | -22.0823
13.3622 | -34.6592 | -22.0906
13.3683 | -34.6597 | -22.0989
13.3744 | -34.6602 | -22.1072
13.3805 | -34.6606 | -22.1158
13.3866 | -34.6611 | -22.124
13.3927 | -34.6616 | -22.1326
13.3989 | -34.662 | -22.1409
13.405 | -34.6625 | -22.1493
13.4111 | -34.663 | -22.1576
13.4172 | -34.6635 | -22.1661
13.4233 | -34.664 | -22.1748
13.4294 | -34.6644 | -22.1833
13.4355 | -34.6649 | -22.1915
13.4416 | -34.6653 | -22.1998
13.4477 | -34.6658 | -22.2082
13.4538 | -34.6663 | -22.2167
13.4599 | -34.6668 | -22.2253
13.466 | -34.6672 | -22.2338
13.4721 | -34.6677 | -22.2422
13.4782 | -34.6682 | -22.2507
13.4843 | -34.6687 | -22.2593
13.4904 | -34.6692 | -22.2675
13.4965 | -34.6696 | -22.2762
13.5026 | -34.6701 | -22.2849
13.5087 | -34.6706 | -22.293
13.5148 | -34.6711 | -22.3017
13.5209 | -34.6715 | -22.3103
13.527 | -34.672 | -22.3188
13.5331 | -34.6725 | -22.3271
13.5393 | -34.673 | -22.3355
13.5454 | -34.6735 | -22.3438
13.5515 | -34.674 | -22.3523
13.5576 | -34.6744 | -22.3607
13.5637 | -34.6749 | -22.3692
13.5698 | -34.6754 | -22.3777
13.5759 | -34.6759 | -22.3857
13.582 | -34.6763 | -22.3941
13.5881 | -34.6768 | -22.4024
13.5942 | -34.6773 | -22.4111
13.6003 | -34.6778 | -22.4194
13.6064 | -34.6783 | -22.428
13.6125 | -34.6788 | -22.4365
13.6186 | -34.6793 | -22.4449
13.6247 | -34.6798 | -22.4536
13.6308 | -34.6803 | -22.462
13.6369 | -34.6807 | -22.4707
13.643 | -34.6812 | -22.4792
13.6491 | -34.6817 | -22.4875
13.6552 | -34.6822 | -22.4959
13.6613 | -34.6826 | -22.5043
13.6674 | -34.6831 | -22.5127
13.6735 | -34.6836 | -22.5211
13.6796 | -34.684 | -22.5295
13.6858 | -34.6845 | -22.5379
13.6919 | -34.685 | -22.5463
13.698 | -34.6855 | -22.5547
13.7041 | -34.686 | -22.5632
13.7102 | -34.6865 | -22.5716
13.7163 | -34.687 | -22.5799
13.7224 | -34.6874 | -22.5882
13.7285 | -34.6879 | -22.5967
13.7346 | -34.6884 | -22.6052
13.7407 | -34.6889 | -22.6136
13.7468 | -34.6894 | -22.622
13.7529 | -34.6898 | -22.6303
13.759 | -34.6903 | -22.6388
13.7651 | -34.6908 | -22.6473
13.7712 | -34.6913 | -22.6558
13.7773 | -34.6917 | -22.6643
13.7834 | -34.6922 | -22.6727
13.7895 | -34.6927 | -22.681
13.7956 | -34.6932 | -22.6895
13.8017 | -34.6937 | -22.6979
13.8078 | -34.6941 | -22.7066
13.8139 | -34.6946 | -22.7147
13.82 | -34.6951 | -22.7229
13.8262 | -34.6956 | -22.7313
13.8323 | -34.6961 | -22.7397
13.8384 | -34.6965 | -22.7482
13.8445 | -34.6971 | -22.7567
13.8506 | -34.6975 | -22.765
13.8567 | -34.698 | -22.7735
13.8628 | -34.6985 | -22.782
13.8689 | -34.699 | -22.7906
13.875 | -34.6994 | -22.799
13.8811 | -34.6999 | -22.8073
13.8872 | -34.7003 | -22.8157
13.8933 | -34.7008 | -22.8237
13.8994 | -34.7013 | -22.8322
13.9055 | -34.7018 | -22.8407
13.9116 | -34.7022 | -22.8491
13.9177 | -34.7027 | -22.8574
13.9238 | -34.7032 | -22.8658
13.9299 | -34.7037 | -22.8739
13.936 | -34.7042 | -22.8824
13.9421 | -34.7047 | -22.8907
13.9482 | -34.7051 | -22.899
13.9543 | -34.7056 | -22.9075
13.9604 | -34.7061 | -22.9159
13.9665 | -34.7066 | -22.924
13.9727 | -34.7071 | -22.9327
13.9788 | -34.7076 | -22.9411
13.9849 | -34.708 | -22.9494
13.991 | -34.7085 | -22.9578
13.9971 | -34.709 | -22.966
14.0032 | -34.7094 | -22.9744
14.0093 | -34.7099 | -22.9827
14.0154 | -34.7104 | -22.9908
14.0215 | -34.7108 | -22.9991
14.0276 | -34.7113 | -23.0075
14.0337 | -34.7118 | -23.0157
14.0398 | -34.7123 | -23.0239
14.0459 | -34.7128 | -23.0321
14.052 | -34.7133 | -23.0407
14.0581 | -34.7137 | -23.049
14.0642 | -34.7142 | -23.0574
14.0703 | -34.7147 | -23.0658
14.0764 | -34.7152 | -23.0739
14.0825 | -34.7156 | -23.0823
14.0886 | -34.7161 | -23.0907
14.0947 | -34.7166 | -23.099
14.1008 | -34.7171 | -23.1075
14.1069 | -34.7176 | -23.1156
14.1131 | -34.718 | -23.124
14.1192 | -34.7185 | -23.1323
14.1253 | -34.719 | -23.1404
14.1314 | -34.7194 | -23.1488
14.1375 | -34.7199 | -23.1571
14.1436 | -34.7204 | -23.1654
14.1497 | -34.7209 | -23.1737
14.1558 | -34.7214 | -23.1822
14.1619 | -34.7218 | -23.1904
14.168 | -34.7223 | -23.1986
14.1741 | -34.7228 | -23.2069
14.1802 | -34.7233 | -23.2151
14.1863 | -34.7237 | -23.2233
14.1924 | -34.7243 | -23.2316
14.1985 | -34.7248 | -23.2399
14.2046 | -34.7253 | -23.2483
14.2107 | -34.7257 | -23.2564
14.2168 | -34.7262 | -23.2649
14.2229 | -34.7266 | -23.2731
14.229 | -34.7271 | -23.2815
14.2351 | -34.7276 | -23.2898
14.2412 | -34.7281 | -23.2981
14.2473 | -34.7286 | -23.3063
14.2534 | -34.729 | -23.3144
14.2596 | -34.7295 | -23.3225
14.2657 | -34.7299 | -23.3308
14.2718 | -34.7304 | -23.339
14.2779 | -34.7309 | -23.3474
14.284 | -34.7314 | -23.3556
14.2901 | -34.7319 | -23.3639
14.2962 | -34.7324 | -23.3722
14.3023 | -34.7328 | -23.3805
14.3084 | -34.7333 | -23.3888
14.3145 | -34.7338 | -23.397
14.3206 | -34.7343 | -23.4054
14.3267 | -34.7348 | -23.4138
14.3328 | -34.7353 | -23.422
14.3389 | -34.7358 | -23.4302
14.345 | -34.7362 | -23.4385
14.3511 | -34.7367 | -23.4467
14.3572 | -34.7372 | -23.4552
14.3633 | -34.7377 | -23.4635
14.3694 | -34.7382 | -23.4717
14.3755 | -34.7387 | -23.4801
14.3816 | -34.7392 | -23.4884
14.3877 | -34.7396 | -23.4964
14.3938 | -34.7401 | -23.5048
14.4 | -34.7406 | -23.513
14.4061 | -34.7411 | -23.5213
14.4122 | -34.7415 | -23.5295
14.4183 | -34.742 | -23.5376
14.4244 | -34.7425 | -23.5457
14.4305 | -34.743 | -23.554
14.4366 | -34.7435 | -23.5623
14.4427 | -34.744 | -23.5707
14.4488 | -34.7445 | -23.579
14.4549 | -34.745 | -23.587
14.461 | -34.7454 | -23.5952
14.4671 | -34.7459 | -23.6036
14.4732 | -34.7464 | -23.6119
14.4793 | -34.7469 | -23.62
14.4854 | -34.7474 | -23.6282
14.4915 | -34.7479 | -23.6361
14.4976 | -34.7483 | -23.6441
14.5037 | -34.7489 | -23.6522
14.5098 | -34.7493 | -23.6606
14.5159 | -34.7498 | -23.6689
14.522 | -34.7502 | -23.677
14.5281 | -34.7507 | -23.685
14.5342 | -34.7512 | -23.6933
14.5403 | -34.7517 | -23.7015
14.5465 | -34.7522 | -23.7097
14.5526 | -34.7527 | -23.7181
14.5587 | -34.7532 | -23.7266
14.5648 | -34.7537 | -23.7346
14.5709 | -34.7542 | -23.7427
14.577 | -34.7546 | -23.7509
14.5831 | -34.7551 | -23.7591
14.5892 | -34.7555 | -23.7675
14.5953 | -34.756 | -23.7757
14.6014 | -34.7565 | -23.7839
14.6075 | -34.757 | -23.792
14.6136 | -34.7575 | -23.8004
14.6197 | -34.7579 | -23.8086
14.6258 | -34.7584 | -23.8169
14.6319 | -34.7589 | -23.8249
14.638 | -34.7594 | -23.833
14.6441 | -34.7599 | -23.8413
14.6502 | -34.7604 | -23.8497
14.6563 | -34.7608 | -23.8579
14.6624 | -34.7613 | -23.8662
14.6685 | -34.7618 | -23.8741
14.6746 | -34.7623 | -23.8824
14.6807 | -34.7628 | -23.8905
14.6869 | -34.7633 | -23.8987
14.693 | -34.7638 | -23.9068
14.6991 | -34.7642 | -23.9149
14.7052 | -34.7648 | -23.9229
14.7113 | -34.7652 | -23.9311
14.7174 | -34.7657 | -23.9393
14.7235 | -34.7662 | -23.9473
14.7296 | -34.7667 | -23.9554
14.7357 | -34.7672 | -23.9634
14.7418 | -34.7676 | -23.9716
14.7479 | -34.7681 | -23.9797
14.754 | -34.7686 | -23.988
14.7601 | -34.7691 | -23.996
14.7662 | -34.7696 | -24.0042
14.7723 | -34.7701 | -24.0121
14.7784 | -34.7705 | -24.0203
14.7845 | -34.771 | -24.0284
14.7906 | -34.7715 | -24.0366
14.7967 | -34.772 | -24.0447
14.8028 | -34.7724 | -24.0528
14.8089 | -34.773 | -24.0609
14.815 | -34.7734 | -24.0691
14.8211 | -34.7739 | -24.0772
14.8272 | -34.7744 | -24.0854
14.8334 | -34.7749 | -24.0935
14.8395 | -34.7754 | -24.1016
14.8456 | -34.7759 | -24.1096
14.8517 | -34.7764 | -24.1178
14.8578 | -34.7769 | -24.126
14.8639 | -34.7773 | -24.1342
14.87 | -34.7778 | -24.1423
14.8761 | -34.7783 | -24.1503
14.8822 | -34.7788 | -24.1586
14.8883 | -34.7793 | -24.1667
14.8944 | -34.7798 | -24.1747
14.9005 | -34.7803 | -24.1828
14.9066 | -34.7807 | -24.191
14.9127 | -34.7812 | -24.199
14.9188 | -34.7817 | -24.2071
14.9249 | -34.7822 | -24.2153
14.931 | -34.7827 | -24.2235
14.9371 | -34.7832 | -24.2316
14.9432 | -34.7837 | -24.2398
14.9493 | -34.7842 | -24.2478
14.9554 | -34.7846 | -24.2559
14.9615 | -34.7851 | -24.264
14.9676 | -34.7856 | -24.2723
14.9738 | -34.7861 | -24.2802
14.9799 | -34.7866 | -24.2884
14.986 | -34.7871 | -24.2962
14.9921 | -34.7875 | -24.3043
14.9982 | -34.788 | -24.3124
15.0043 | -34.7885 | -24.3205
15.0104 | -34.789 | -24.3287
15.0165 | -34.7895 | -24.3369
15.0226 | -34.79 | -24.3449
15.0287 | -34.7905 | -24.353
15.0348 | -34.7909 | -24.3611
15.0409 | -34.7914 | -24.3692
15.047 | -34.7919 | -24.3774
15.0531 | -34.7925 | -24.3854
15.0592 | -34.7929 | -24.3935
15.0653 | -34.7934 | -24.4016
15.0714 | -34.7939 | -24.4095
15.0775 | -34.7944 | -24.4175
15.0836 | -34.7949 | -24.4256
15.0897 | -34.7954 | -24.4335
15.0958 | -34.7959 | -24.4415
15.1019 | -34.7964 | -24.4496
15.108 | -34.7969 | -24.4576
15.1141 | -34.7974 | -24.4656
15.1203 | -34.7979 | -24.4738
15.1264 | -34.7984 | -24.4819
15.1325 | -34.7989 | -24.4899
15.1386 | -34.7994 | -24.4978
15.1447 | -34.7999 | -24.5059
15.1508 | -34.8004 | -24.514
15.1569 | -34.8009 | -24.522
15.163 | -34.8014 | -24.5298
15.1691 | -34.8019 | -24.5378
15.1752 | -34.8024 | -24.5459
15.1813 | -34.8029 | -24.5538
15.1874 | -34.8034 | -24.5619
15.1935 | -34.8039 | -24.5697
15.1996 | -34.8044 | -24.5776
15.2057 | -34.8049 | -24.5856
15.2118 | -34.8054 | -24.5937
15.2179 | -34.8059 | -24.6019
15.224 | -34.8064 | -24.6099
15.2301 | -34.8069 | -24.6178
15.2362 | -34.8074 | -24.6258
15.2423 | -34.8079 | -24.634
15.2484 | -34.8084 | -24.6422
15.2545 | -34.8089 | -24.6503
15.2607 | -34.8094 | -24.6583
15.2668 | -34.8099 | -24.6661
15.2729 | -34.8104 | -24.6739
15.279 | -34.8109 | -24.6822
15.2851 | -34.8114 | -24.6901
15.2912 | -34.8119 | -24.6982
15.2973 | -34.8124 | -24.7064
15.3034 | -34.8129 | -24.7143
15.3095 | -34.8134 | -24.7224
15.3156 | -34.8139 | -24.7303
15.3217 | -34.8144 | -24.7382
15.3278 | -34.8149 | -24.7463
15.3339 | -34.8154 | -24.7541
15.34 | -34.8159 | -24.7619
15.3461 | -34.8164 | -24.7699
15.3522 | -34.8169 | -24.7779
15.3583 | -34.8174 | -24.7861
15.3644 | -34.8179 | -24.794
15.3705 | -34.8184 | -24.8017
15.3766 | -34.8189 | -24.8097
15.3827 | -34.8194 | -24.8175
15.3888 | -34.8199 | -24.8255
15.3949 | -34.8204 | -24.8336
15.401 | -34.8209 | -24.8417
15.4072 | -34.8214 | -24.8494
15.4133 | -34.8218 | -24.8574
15.4194 | -34.8223 | -24.8654
15.4255 | -34.8228 | -24.8735
15.4316 | -34.8233 | -24.8814
15.4377 | -34.8238 | -24.8895
15.4438 | -34.8243 | -24.8974
15.4499 | -34.8248 | -24.9053
15.456 | -34.8253 | -24.9132
15.4621 | -34.8258 | -24.9212
15.4682 | -34.8263 | -24.9295
15.4743 | -34.8268 | -24.9373
15.4804 | -34.8273 | -24.9452
15.4865 | -34.8278 | -24.9531
15.4926 | -34.8283 | -24.9613
15.4987 | -34.8288 | -24.9692
15.5048 | -34.8293 | -24.9773
15.5109 | -34.8298 | -24.985
15.517 | -34.8304 | -24.993
15.5231 | -34.8309 | -25.001
15.5292 | -34.8313 | -25.0089
15.5353 | -34.8318 | -25.0168
15.5414 | -34.8323 | -25.0247
15.5476 | -34.8329 | -25.0322
15.5537 | -34.8334 | -25.0401
15.5598 | -34.8339 | -25.0482
15.5659 | -34.8344 | -25.0561
15.572 | -34.8349 | -25.0641
15.5781 | -34.8354 | -25.0719
15.5842 | -34.8359 | -25.0799
15.5903 | -34.8364 | -25.0876
15.5964 | -34.8369 | -25.0955
15.6025 | -34.8374 | -25.1035
15.6086 | -34.8379 | -25.1115
15.6147 | -34.8384 | -25.1195
15.6208 | -34.8389 | -25.1271
15.6269 | -34.8394 | -25.135
15.633 | -34.8399 | -25.143
15.6391 | -34.8404 | -25.151
15.6452 | -34.8409 | -25.1589
15.6513 | -34.8414 | -25.1666
15.6574 | -34.8419 | -25.1747
15.6635 | -34.8425 | -25.1825
15.6696 | -34.843 | -25.1905
15.6757 | -34.8434 | -25.1985
15.6818 | -34.8439 | -25.2063
15.688 | -34.8444 | -25.2142
15.6941 | -34.8449 | -25.2221
15.7002 | -34.8454 | -25.2299
15.7063 | -34.8459 | -25.2377
15.7124 | -34.8464 | -25.2457
15.7185 | -34.8469 | -25.2536
15.7246 | -34.8474 | -25.2614
15.7307 | -34.8479 | -25.2692
15.7368 | -34.8484 | -25.277
15.7429 | -34.8489 | -25.2848
15.749 | -34.8495 | -25.2928
15.7551 | -34.85 | -25.3007
15.7612 | -34.8505 | -25.3084
15.7673 | -34.851 | -25.3163
15.7734 | -34.8515 | -25.3241
15.7795 | -34.852 | -25.3321
15.7856 | -34.8525 | -25.3398
15.7917 | -34.853 | -25.3476
15.7978 | -34.8535 | -25.3556
15.8039 | -34.854 | -25.3634
15.81 | -34.8545 | -25.3713
15.8161 | -34.855 | -25.3792
15.8222 | -34.8555 | -25.387
15.8283 | -34.8561 | -25.3946
15.8345 | -34.8566 | -25.4023
15.8406 | -34.857 | -25.4102
15.8467 | -34.8575 | -25.4181
15.8528 | -34.8581 | -25.4259
15.8589 | -34.8586 | -25.4336
15.865 | -34.8591 | -25.4412
15.8711 | -34.8596 | -25.4491
15.8772 | -34.8601 | -25.4568
15.8833 | -34.8606 | -25.4648
15.8894 | -34.8611 | -25.4728
15.8955 | -34.8616 | -25.4806
15.9016 | -34.8621 | -25.4882
15.9077 | -34.8626 | -25.496
15.9138 | -34.8631 | -25.5039
15.9199 | -34.8636 | -25.5116
15.926 | -34.8641 | -25.5193
15.9321 | -34.8646 | -25.527
15.9382 | -34.8651 | -25.5347
15.9443 | -34.8656 | -25.5425
15.9504 | -34.8661 | -25.5504
15.9565 | -34.8666 | -25.5583
15.9626 | -34.867 | -25.5662
15.9687 | -34.8675 | -25.574
15.9749 | -34.868 | -25.5818
15.981 | -34.8685 | -25.5897
15.9871 | -34.869 | -25.5975
15.9932 | -34.8695 | -25.6051
15.9993 | -34.87 | -25.6129
16.0054 | -34.8705 | -25.6205
16.0115 | -34.871 | -25.6282
16.0176 | -34.8715 | -25.6361
16.0237 | -34.872 | -25.6438
16.0298 | -34.8725 | -25.6515
16.0359 | -34.873 | -25.6593
16.042 | -34.8735 | -25.6671
16.0481 | -34.874 | -25.6748
16.0542 | -34.8745 | -25.6828
16.0603 | -34.875 | -25.6905
16.0664 | -34.8755 | -25.6983
16.0725 | -34.876 | -25.7059
16.0786 | -34.8765 | -25.7137
16.0847 | -34.877 | -25.7215
16.0908 | -34.8774 | -25.7294
16.0969 | -34.8779 | -25.7373
16.103 | -34.8784 | -25.7451
16.1091 | -34.879 | -25.7529
16.1152 | -34.8795 | -25.7606
16.1214 | -34.88 | -25.7683
16.1275 | -34.8805 | -25.776
16.1336 | -34.881 | -25.7837
16.1397 | -34.8815 | -25.7912
16.1458 | -34.882 | -25.7987
16.1519 | -34.8825 | -25.8065
16.158 | -34.883 | -25.8146
16.1641 | -34.8835 | -25.8225
16.1702 | -34.884 | -25.8303
16.1763 | -34.8845 | -25.8381
16.1824 | -34.885 | -25.8457
16.1885 | -34.8855 | -25.8533
16.1946 | -34.886 | -25.861
16.2007 | -34.8865 | -25.8686
16.2068 | -34.887 | -25.8764
16.2129 | -34.8875 | -25.8842
16.219 | -34.888 | -25.8917
16.2251 | -34.8885 | -25.8993
16.2312 | -34.889 | -25.9072
16.2373 | -34.8895 | -25.9149
16.2434 | -34.89 | -25.9225
16.2495 | -34.8905 | -25.9299
16.2556 | -34.891 | -25.938
16.2618 | -34.8915 | -25.9457
16.2679 | -34.892 | -25.9535
16.274 | -34.8925 | -25.9613
16.2801 | -34.893 | -25.969
16.2862 | -34.8935 | -25.9764
16.2923 | -34.894 | -25.9841
16.2984 | -34.8945 | -25.992
16.3045 | -34.895 | -25.9997
16.3106 | -34.8955 | -26.0076
16.3167 | -34.896 | -26.0153
16.3228 | -34.8965 | -26.0228
16.3289 | -34.897 | -26.0305
16.335 | -34.8975 | -26.038
16.3411 | -34.898 | -26.0456
16.3472 | -34.8985 | -26.0534
16.3533 | -34.899 | -26.0612
16.3594 | -34.8995 | -26.0686
16.3655 | -34.9 | -26.0762
16.3716 | -34.9005 | -26.0838
16.3777 | -34.901 | -26.0914
16.3838 | -34.9015 | -26.0989
16.3899 | -34.902 | -26.1066
16.396 | -34.9025 | -26.1144
16.4021 | -34.903 | -26.122
16.4083 | -34.9035 | -26.1295
16.4144 | -34.904 | -26.1372
16.4205 | -34.9045 | -26.1448
16.4266 | -34.9051 | -26.1523
16.4327 | -34.9056 | -26.1599
16.4388 | -34.9061 | -26.1676
16.4449 | -34.9066 | -26.1752
16.451 | -34.9071 | -26.183
16.4571 | -34.9076 | -26.1908
16.4632 | -34.9081 | -26.1983
16.4693 | -34.9086 | -26.206
16.4754 | -34.9091 | -26.2137
16.4815 | -34.9096 | -26.2214
16.4876 | -34.9101 | -26.229
16.4937 | -34.9105 | -26.2366
16.4998 | -34.9111 | -26.2439
16.5059 | -34.9116 | -26.2514
16.512 | -34.9121 | -26.2591
16.5181 | -34.9126 | -26.2669
16.5242 | -34.9131 | -26.2747
16.5303 | -34.9136 | -26.2824
16.5364 | -34.9141 | -26.29
16.5425 | -34.9146 | -26.2976
16.5487 | -34.915 | -26.3052
16.5548 | -34.9155 | -26.313
16.5609 | -34.916 | -26.3205
16.567 | -34.9165 | -26.328
16.5731 | -34.9171 | -26.3358
16.5792 | -34.9176 | -26.3434
16.5853 | -34.9181 | -26.351
16.5914 | -34.9186 | -26.3585
16.5975 | -34.9191 | -26.3663
16.6036 | -34.9196 | -26.3738
16.6097 | -34.9201 | -26.3815
16.6158 | -34.9206 | -26.3891
16.6219 | -34.9211 | -26.3966
16.628 | -34.9216 | -26.4043
16.6341 | -34.9221 | -26.412
16.6402 | -34.9226 | -26.4194
16.6463 | -34.9231 | -26.4274
16.6524 | -34.9236 | -26.4349
16.6585 | -34.9241 | -26.4425
16.6646 | -34.9246 | -26.4499
16.6707 | -34.9251 | -26.4574
16.6768 | -34.9256 | -26.465
16.6829 | -34.9261 | -26.4725
16.689 | -34.9266 | -26.4801
16.6952 | -34.9271 | -26.4877
16.7013 | -34.9276 | -26.4951
16.7074 | -34.9281 | -26.5025
16.7135 | -34.9286 | -26.5101
16.7196 | -34.9291 | -26.5177
16.7257 | -34.9296 | -26.5252
16.7318 | -34.9301 | -26.5327
16.7379 | -34.9306 | -26.5403
16.744 | -34.931 | -26.5476
16.7501 | -34.9315 | -26.5552
16.7562 | -34.9321 | -26.5626
16.7623 | -34.9325 | -26.57
16.7684 | -34.933 | -26.5777
16.7745 | -34.9335 | -26.5852
16.7806 | -34.934 | -26.5924
16.7867 | -34.9346 | -26.6
16.7928 | -34.9351 | -26.6074
16.7989 | -34.9356 | -26.615
16.805 | -34.936 | -26.6226
16.8111 | -34.9365 | -26.6299
16.8172 | -34.937 | -26.6374
16.8233 | -34.9375 | -26.645
16.8294 | -34.938 | -26.6526
16.8356 | -34.9386 | -26.6603
16.8417 | -34.9391 | -26.6678
16.8478 | -34.9396 | -26.6753
16.8539 | -34.9401 | -26.6826
16.86 | -34.9406 | -26.6901
16.8661 | -34.9412 | -26.6977
16.8722 | -34.9416 | -26.7052
16.8783 | -34.9421 | -26.7128
16.8844 | -34.9427 | -26.7203
16.8905 | -34.9432 | -26.728
16.8966 | -34.9437 | -26.7355
16.9027 | -34.9442 | -26.7431
16.9088 | -34.9447 | -26.7506
16.9149 | -34.9451 | -26.7582
16.921 | -34.9457 | -26.7655
16.9271 | -34.9462 | -26.7731
16.9332 | -34.9467 | -26.7807
16.9393 | -34.9471 | -26.7883
16.9454 | -34.9476 | -26.7958
16.9515 | -34.9481 | -26.8035
16.9576 | -34.9487 | -26.8108
16.9637 | -34.9492 | -26.818
16.9698 | -34.9497 | -26.8255
16.9759 | -34.9502 | -26.8331
16.9821 | -34.9508 | -26.8411
16.9882 | -34.9513 | -26.8485
16.9943 | -34.9518 | -26.8559
17.0004 | -34.9523 | -26.8636
17.0065 | -34.9528 | -26.871
17.0126 | -34.9533 | -26.8786
17.0187 | -34.9538 | -26.8863
17.0248 | -34.9543 | -26.8937
17.0309 | -34.9548 | -26.901
17.037 | -34.9553 | -26.9087
17.0431 | -34.9558 | -26.9163
17.0492 | -34.9563 | -26.9238
17.0553 | -34.9568 | -26.9313
17.0614 | -34.9573 | -26.9386
17.0675 | -34.9578 | -26.9463
17.0736 | -34.9584 | -26.954
17.0797 | -34.9589 | -26.9612
17.0858 | -34.9594 | -26.9687
17.0919 | -34.9599 | -26.9764
17.098 | -34.9604 | -26.9837
17.1041 | -34.9609 | -26.9913
17.1102 | -34.9614 | -26.9987
17.1163 | -34.9619 | -27.0062
17.1225 | -34.9624 | -27.0135
17.1286 | -34.9629 | -27.0209
17.1347 | -34.9634 | -27.0285
17.1408 | -34.964 | -27.0362
17.1469 | -34.9645 | -27.0438
17.153 | -34.965 | -27.0511
17.1591 | -34.9654 | -27.0587
17.1652 | -34.9659 | -27.0661
17.1713 | -34.9664 | -27.0735
17.1774 | -34.9669 | -27.081
17.1835 | -34.9674 | -27.0883
17.1896 | -34.9679 | -27.0958
17.1957 | -34.9684 | -27.1032
17.2018 | -34.9689 | -27.1103
17.2079 | -34.9694 | -27.1176
17.214 | -34.9699 | -27.1253
17.2201 | -34.9704 | -27.1329
17.2262 | -34.9709 | -27.1406
17.2323 | -34.9714 | -27.1479
17.2384 | -34.9719 | -27.1553
17.2445 | -34.9724 | -27.1628
17.2506 | -34.9729 | -27.1703
17.2567 | -34.9735 | -27.1777
17.2628 | -34.974 | -27.185
17.269 | -34.9745 | -27.1923
17.2751 | -34.975 | -27.1995
17.2812 | -34.9755 | -27.2071
17.2873 | -34.9759 | -27.2146
17.2934 | -34.9764 | -27.2223
17.2995 | -34.977 | -27.2298
17.3056 | -34.9775 | -27.2371
17.3117 | -34.978 | -27.2448
17.3178 | -34.9785 | -27.2522
17.3239 | -34.979 | -27.2594
17.33 | -34.9795 | -27.2668
17.3361 | -34.98 | -27.2743
17.3422 | -34.9805 | -27.2813
17.3483 | -34.9809 | -27.2888
17.3544 | -34.9815 | -27.2964
17.3605 | -34.982 | -27.3039
17.3666 | -34.9825 | -27.3112
17.3727 | -34.983 | -27.3188
17.3788 | -34.9835 | -27.326
17.3849 | -34.984 | -27.3334
17.391 | -34.9845 | -27.3408
17.3971 | -34.985 | -27.3482
17.4032 | -34.9855 | -27.3557
17.4094 | -34.986 | -27.3631
17.4155 | -34.9865 | -27.3705
17.4216 | -34.987 | -27.378
17.4277 | -34.9876 | -27.3853
17.4338 | -34.9881 | -27.3928
17.4399 | -34.9886 | -27.4002
17.446 | -34.9891 | -27.4072
17.4521 | -34.9896 | -27.4144
17.4582 | -34.9901 | -27.4218
17.4643 | -34.9906 | -27.4291
17.4704 | -34.9911 | -27.4362
17.4765 | -34.9916 | -27.4435
17.4826 | -34.9921 | -27.4508
17.4887 | -34.9926 | -27.4581
17.4948 | -34.9931 | -27.4656
17.5009 | -34.9936 | -27.4726
17.507 | -34.9941 | -27.4799
17.5131 | -34.9946 | -27.4872
17.5192 | -34.9951 | -27.4942
17.5253 | -34.9955 | -27.5016
17.5314 | -34.9961 | -27.5091
17.5375 | -34.9966 | -27.5166
17.5436 | -34.997 | -27.5239
17.5497 | -34.9975 | -27.5313
17.5559 | -34.998 | -27.5383
17.562 | -34.9986 | -27.5456
17.5681 | -34.9991 | -27.5531
17.5742 | -34.9996 | -27.5606
17.5803 | -35.0001 | -27.568
17.5864 | -35.0007 | -27.5752
17.5925 | -35.0012 | -27.5828
17.5986 | -35.0017 | -27.59
17.6047 | -35.0022 | -27.5975
17.6108 | -35.0027 | -27.6046
17.6169 | -35.0032 | -27.6118
17.623 | -35.0037 | -27.619
17.6291 | -35.0042 | -27.6264
17.6352 | -35.0047 | -27.6339
17.6413 | -35.0052 | -27.6412
17.6474 | -35.0057 | -27.6486
17.6535 | -35.0062 | -27.6557
17.6596 | -35.0068 | -27.6627
17.6657 | -35.0073 | -27.6702
17.6718 | -35.0078 | -27.6776
17.6779 | -35.0083 | -27.685
17.684 | -35.0088 | -27.6924
17.6901 | -35.0093 | -27.6998
17.6963 | -35.0098 | -27.7069
17.7024 | -35.0103 | -27.7141
17.7085 | -35.0108 | -27.7212
17.7146 | -35.0113 | -27.7283
17.7207 | -35.0118 | -27.7357
17.7268 | -35.0123 | -27.7428
17.7329 | -35.0128 | -27.7502
17.739 | -35.0133 | -27.7576
17.7451 | -35.0138 | -27.7648
17.7512 | -35.0143 | -27.7721
17.7573 | -35.0148 | -27.7796
17.7634 | -35.0153 | -27.7866
17.7695 | -35.0159 | -27.794
17.7756 | -35.0164 | -27.8013
17.7817 | -35.0169 | -27.8087
17.7878 | -35.0174 | -27.8162
17.7939 | -35.0179 | -27.8233
17.8 | -35.0184 | -27.8305
17.8061 | -35.0189 | -27.8378
17.8122 | -35.0194 | -27.8452
17.8183 | -35.0199 | -27.8526
17.8244 | -35.0204 | -27.8599
17.8305 | -35.0209 | -27.8671
17.8366 | -35.0214 | -27.8743
17.8428 | -35.0219 | -27.8817
17.8489 | -35.0224 | -27.8889
17.855 | -35.0229 | -27.8961
17.8611 | -35.0235 | -27.9032
17.8672 | -35.024 | -27.9102
17.8733 | -35.0245 | -27.9175
17.8794 | -35.025 | -27.9248
17.8855 | -35.0255 | -27.9322
17.8916 | -35.0259 | -27.9398
17.8977 | -35.0264 | -27.9472
17.9038 | -35.027 | -27.9543
17.9099 | -35.0275 | -27.9616
17.916 | -35.028 | -27.9688
17.9221 | -35.0285 | -27.9761
17.9282 | -35.029 | -27.9833
17.9343 | -35.0295 | -27.9907
17.9404 | -35.03 | -27.9977
17.9465 | -35.0305 | -28.005
17.9526 | -35.031 | -28.0123
17.9587 | -35.0315 | -28.0195
17.9648 | -35.032 | -28.0266
17.9709 | -35.0325 | -28.034
17.977 | -35.033 | -28.0412
17.9832 | -35.0335 | -28.0483
17.9893 | -35.034 | -28.0557
17.9954 | -35.0345 | -28.0632
18.0015 | -35.035 | -28.0707
18.0076 | -35.0355 | -28.0776
18.0137 | -35.036 | -28.0848
18.0198 | -35.0365 | -28.092
18.0259 | -35.0369 | -28.0991
18.032 | -35.0374 | -28.1063
18.0381 | -35.0379 | -28.1136
18.0442 | -35.0384 | -28.1206
18.0503 | -35.0389 | -28.128
18.0564 | -35.0394 | -28.1354
18.0625 | -35.0399 | -28.1425
18.0686 | -35.0404 | -28.1499
18.0747 | -35.0409 | -28.1572
18.0808 | -35.0415 | -28.1643
18.0869 | -35.042 | -28.1715
18.093 | -35.0425 | -28.1786
18.0991 | -35.043 | -28.1859
18.1052 | -35.0435 | -28.1931
18.1113 | -35.044 | -28.2006
18.1174 | -35.0445 | -28.2077
18.1236 | -35.045 | -28.2148
18.1297 | -35.0455 | -28.222
18.1358 | -35.046 | -28.2294
18.1419 | -35.0465 | -28.2365
18.148 | -35.047 | -28.2438
18.1541 | -35.0476 | -28.2511
18.1602 | -35.0481 | -28.2586
18.1663 | -35.0486 | -28.2658
18.1724 | -35.0491 | -28.273
18.1785 | -35.0495 | -28.2801
18.1846 | -35.05 | -28.2873
18.1907 | -35.0505 | -28.2944
18.1968 | -35.051 | -28.3019
18.2029 | -35.0515 | -28.3092
18.209 | -35.052 | -28.3165
18.2151 | -35.0525 | -28.3236
18.2212 | -35.053 | -28.3306
18.2273 | -35.0535 | -28.3379
18.2334 | -35.0541 | -28.345
18.2395 | -35.0546 | -28.3521
18.2456 | -35.0551 | -28.3592
18.2517 | -35.0556 | -28.3664
18.2578 | -35.0561 | -28.3733
18.2639 | -35.0566 | -28.3805
18.2701 | -35.0571 | -28.3878
18.2762 | -35.0576 | -28.3949
18.2823 | -35.0581 | -28.4018
18.2884 | -35.0586 | -28.4086
18.2945 | -35.0591 | -28.4154
18.3006 | -35.0596 | -28.4226
18.3067 | -35.0601 | -28.4298
18.3128 | -35.0607 | -28.4371
18.3189 | -35.0612 | -28.4443
18.325 | -35.0617 | -28.4513
18.3311 | -35.0622 | -28.4586
18.3372 | -35.0627 | -28.4659
18.3433 | -35.0632 | -28.4733
18.3494 | -35.0637 | -28.4804
18.3555 | -35.0642 | -28.4876
18.3616 | -35.0647 | -28.4944
18.3677 | -35.0652 | -28.5017
18.3738 | -35.0657 | -28.5088
18.3799 | -35.0662 | -28.5162
18.386 | -35.0667 | -28.5234
18.3921 | -35.0672 | -28.5306
18.3982 | -35.0677 | -28.5374
18.4043 | -35.0682 | -28.5447
18.4105 | -35.0687 | -28.5519
18.4166 | -35.0692 | -28.5588
18.4227 | -35.0697 | -28.566
18.4288 | -35.0702 | -28.5732
18.4349 | -35.0707 | -28.58
18.441 | -35.0712 | -28.5871
18.4471 | -35.0718 | -28.5941
18.4532 | -35.0723 | -28.6013
18.4593 | -35.0728 | -28.6084
18.4654 | -35.0733 | -28.6154
18.4715 | -35.0738 | -28.6228
18.4776 | -35.0743 | -28.6301
18.4837 | -35.0748 | -28.6373
18.4898 | -35.0753 | -28.6445
18.4959 | -35.0758 | -28.6518
18.502 | -35.0762 | -28.659
18.5081 | -35.0767 | -28.6663
18.5142 | -35.0772 | -28.6734
18.5203 | -35.0777 | -28.6805
18.5264 | -35.0782 | -28.6876
18.5325 | -35.0787 | -28.6948
18.5386 | -35.0793 | -28.7019
18.5447 | -35.0798 | -28.7091
18.5508 | -35.0803 | -28.7159
18.557 | -35.0808 | -28.7232
18.5631 | -35.0813 | -28.7303
18.5692 | -35.0818 | -28.7374
18.5753 | -35.0823 | -28.7446
18.5814 | -35.0828 | -28.7517
18.5875 | -35.0834 | -28.7588
18.5936 | -35.0839 | -28.7659
18.5997 | -35.0843 | -28.7731
18.6058 | -35.0848 | -28.78
18.6119 | -35.0853 | -28.7872
18.618 | -35.0858 | -28.7944
18.6241 | -35.0863 | -28.8014
18.6302 | -35.0867 | -28.8085
18.6363 | -35.0872 | -28.8157
18.6424 | -35.0878 | -28.8228
18.6485 | -35.0882 | -28.8299
18.6546 | -35.0887 | -28.8372
18.6607 | -35.0892 | -28.8442
18.6668 | -35.0897 | -28.8515
18.6729 | -35.0903 | -28.8587
18.679 | -35.0908 | -28.8655
18.6851 | -35.0913 | -28.8726
18.6912 | -35.0918 | -28.8795
18.6974 | -35.0924 | -28.8866
18.7035 | -35.0929 | -28.8939
18.7096 | -35.0934 | -28.9011
18.7157 | -35.0939 | -28.9083
18.7218 | -35.0944 | -28.9153
18.7279 | -35.0949 | -28.9224
18.734 | -35.0954 | -28.9296
18.7401 | -35.096 | -28.9367
18.7462 | -35.0964 | -28.9436
18.7523 | -35.0969 | -28.9508
18.7584 | -35.0975 | -28.958
18.7645 | -35.098 | -28.965
18.7706 | -35.0984 | -28.972
18.7767 | -35.0989 | -28.9791
18.7828 | -35.0994 | -28.9859
18.7889 | -35.0999 | -28.993
18.795 | -35.1004 | -29.0004
18.8011 | -35.1009 | -29.0077
18.8072 | -35.1014 | -29.015
18.8133 | -35.1019 | -29.0221
18.8194 | -35.1024 | -29.029
18.8255 | -35.1029 | -29.0362
18.8316 | -35.1034 | -29.0435
18.8377 | -35.1039 | -29.0506
18.8439 | -35.1044 | -29.0579
18.85 | -35.1049 | -29.0649
18.8561 | -35.1054 | -29.0718
18.8622 | -35.1059 | -29.0789
18.8683 | -35.1064 | -29.0858
18.8744 | -35.1069 | -29.093
18.8805 | -35.1074 | -29.1003
18.8866 | -35.1079 | -29.1075
18.8927 | -35.1084 | -29.1144
18.8988 | -35.109 | -29.1213
18.9049 | -35.1095 | -29.1284
18.911 | -35.11 | -29.1357
18.9171 | -35.1105 | -29.1431
18.9232 | -35.111 | -29.15
18.9293 | -35.1115 | -29.1571
18.9354 | -35.112 | -29.1641
18.9415 | -35.1125 | -29.1711
18.9476 | -35.113 | -29.178
18.9537 | -35.1135 | -29.1851
18.9598 | -35.1141 | -29.1921
18.9659 | -35.1145 | -29.1991
18.972 | -35.115 | -29.2063
18.9781 | -35.1155 | -29.2135
18.9843 | -35.116 | -29.2208
18.9904 | -35.1165 | -29.2278
18.9965 | -35.117 | -29.2346
19.0026 | -35.1175 | -29.2418
19.0087 | -35.1181 | -29.2489
19.0148 | -35.1186 | -29.2557
19.0209 | -35.119 | -29.2629
19.027 | -35.1195 | -29.27
19.0331 | -35.12 | -29.2771
19.0392 | -35.1205 | -29.2841
19.0453 | -35.1211 | -29.2913
19.0514 | -35.1216 | -29.2983
19.0575 | -35.1221 | -29.3053
19.0636 | -35.1226 | -29.3121
19.0697 | -35.1231 | -29.319
19.0758 | -35.1236 | -29.326
19.0819 | -35.1241 | -29.3332
19.088 | -35.1246 | -29.3402
19.0941 | -35.1251 | -29.3472
19.1002 | -35.1256 | -29.3543
19.1063 | -35.1261 | -29.3616
19.1124 | -35.1266 | -29.3685
19.1185 | -35.1271 | -29.3751
19.1246 | -35.1276 | -29.3821
19.1308 | -35.1281 | -29.3891
19.1369 | -35.1287 | -29.3959
19.143 | -35.1292 | -29.4031
19.1491 | -35.1296 | -29.4103
19.1552 | -35.1301 | -29.4173
19.1613 | -35.1306 | -29.4243
19.1674 | -35.1311 | -29.4312
19.1735 | -35.1316 | -29.4385
19.1796 | -35.1321 | -29.4454
19.1857 | -35.1326 | -29.4525
19.1918 | -35.1332 | -29.4596
19.1979 | -35.1336 | -29.4666
19.204 | -35.1341 | -29.4735
19.2101 | -35.1346 | -29.4805
19.2162 | -35.1351 | -29.4875
19.2223 | -35.1356 | -29.4947
19.2284 | -35.1361 | -29.502
19.2345 | -35.1366 | -29.5089
19.2406 | -35.1371 | -29.5156
19.2467 | -35.1376 | -29.5223
19.2528 | -35.1381 | -29.5293
19.2589 | -35.1387 | -29.5363
19.265 | -35.1392 | -29.5435
19.2712 | -35.1397 | -29.5503
19.2773 | -35.1402 | -29.557
19.2834 | -35.1407 | -29.5639
19.2895 | -35.1413 | -29.571
19.2956 | -35.1417 | -29.5781
19.3017 | -35.1422 | -29.5851
19.3078 | -35.1427 | -29.5922
19.3139 | -35.1431 | -29.599
19.32 | -35.1436 | -29.6058
19.3261 | -35.1441 | -29.613
19.3322 | -35.1446 | -29.6202
19.3383 | -35.1451 | -29.6269
19.3444 | -35.1455 | -29.634
19.3505 | -35.1461 | -29.6409
19.3566 | -35.1465 | -29.6481
19.3627 | -35.147 | -29.6549
19.3688 | -35.1475 | -29.6619
19.3749 | -35.148 | -29.6687
19.381 | -35.1485 | -29.6752
19.3871 | -35.149 | -29.6823
19.3932 | -35.1495 | -29.6894
19.3993 | -35.15 | -29.6963
19.4054 | -35.1505 | -29.7036
19.4115 | -35.151 | -29.7104
19.4177 | -35.1516 | -29.7174
19.4238 | -35.1521 | -29.7244
19.4299 | -35.1526 | -29.7314
19.436 | -35.1531 | -29.7385
19.4421 | -35.1535 | -29.7456
19.4482 | -35.154 | -29.7527
19.4543 | -35.1545 | -29.7598
19.4604 | -35.1551 | -29.7668
19.4665 | -35.1555 | -29.7741
19.4726 | -35.156 | -29.7812
19.4787 | -35.1565 | -29.7884
19.4848 | -35.157 | -29.7949
19.4909 | -35.1574 | -29.8014
19.497 | -35.158 | -29.808
19.5031 | -35.1585 | -29.8153
19.5092 | -35.159 | -29.8221
19.5153 | -35.1594 | -29.8289
19.5214 | -35.1599 | -29.8358
19.5275 | -35.1604 | -29.8424
19.5336 | -35.1608 | -29.8491
19.5397 | -35.1613 | -29.8559
19.5458 | -35.1618 | -29.8624
19.5519 | -35.1622 | -29.8691
19.5581 | -35.1627 | -29.8762
19.5642 | -35.1632 | -29.883
19.5703 | -35.1638 | -29.8898
19.5764 | -35.1643 | -29.8972
19.5825 | -35.1648 | -29.9041
19.5886 | -35.1653 | -29.9111
19.5947 | -35.1657 | -29.9184
19.6008 | -35.1662 | -29.9252
19.6069 | -35.1668 | -29.9319
19.613 | -35.1673 | -29.9385
19.6191 | -35.1678 | -29.9451
19.6252 | -35.1683 | -29.9516
19.6313 | -35.1688 | -29.9586
19.6374 | -35.1693 | -29.9659
19.6435 | -35.1697 | -29.9726
19.6496 | -35.1703 | -29.9794
19.6557 | -35.1707 | -29.9864
19.6618 | -35.1713 | -29.9936
19.6679 | -35.1718 | -30.0004
19.674 | -35.1723 | -30.0074
19.6801 | -35.1728 | -30.0143
19.6862 | -35.1733 | -30.021
19.6923 | -35.1738 | -30.0276
19.6984 | -35.1744 | -30.0348
19.7046 | -35.175 | -30.0415
19.7107 | -35.1754 | -30.0483
19.7168 | -35.176 | -30.0549
19.7229 | -35.1765 | -30.0616
19.729 | -35.177 | -30.0683
19.7351 | -35.1774 | -30.0757
19.7412 | -35.1778 | -30.0832
19.7473 | -35.1783 | -30.0901
19.7534 | -35.1788 | -30.0973
19.7595 | -35.1793 | -30.1043
19.7656 | -35.1797 | -30.1113
19.7717 | -35.1802 | -30.1174
19.7778 | -35.1807 | -30.1244
19.7839 | -35.1811 | -30.131
19.79 | -35.1816 | -30.138
19.7961 | -35.182 | -30.1444
19.8022 | -35.1825 | -30.1513
19.8083 | -35.183 | -30.1582
19.8144 | -35.1834 | -30.1652
19.8205 | -35.1839 | -30.1721
19.8266 | -35.1844 | -30.1788
19.8327 | -35.1849 | -30.1854
19.8388 | -35.1854 | -30.192
19.845 | -35.1859 | -30.1989
19.8511 | -35.1864 | -30.2059
19.8572 | -35.1868 | -30.2126
19.8633 | -35.1873 | -30.2192
19.8694 | -35.1877 | -30.2265
19.8755 | -35.1884 | -30.2334
19.8816 | -35.1888 | -30.2402
19.8877 | -35.1893 | -30.2469
19.8938 | -35.1898 | -30.2537
19.8999 | -35.1903 | -30.2605
19.906 | -35.1909 | -30.2676
19.9121 | -35.1914 | -30.2744
19.9182 | -35.1918 | -30.2812
19.9243 | -35.1924 | -30.2882
19.9304 | -35.1929 | -30.2954
19.9365 | -35.1934 | -30.3019
19.9426 | -35.1939 | -30.3084
19.9487 | -35.1944 | -30.3155
19.9548 | -35.1949 | -30.3227
19.9609 | -35.1954 | -30.3295
19.967 | -35.1959 | -30.3362
19.9731 | -35.1964 | -30.3434
19.9792 | -35.1969 | -30.3499
19.9853 | -35.1974 | -30.3564
19.9915 | -35.1979 | -30.3631
19.9976 | -35.1983 | -30.3706
20.0037 | -35.1989 | -30.3776
20.0098 | -35.1994 | -30.3843
20.0159 | -35.1998 | -30.3914
20.022 | -35.2003 | -30.3982
20.0281 | -35.2008 | -30.4053
20.0342 | -35.2013 | -30.4121
20.0403 | -35.2019 | -30.4191
20.0464 | -35.2024 | -30.426
20.0525 | -35.2029 | -30.4329
20.0586 | -35.2034 | -30.4399
20.0647 | -35.2039 | -30.4468
20.0708 | -35.2044 | -30.4536
20.0769 | -35.2049 | -30.4604
20.083 | -35.2054 | -30.4667
20.0891 | -35.2059 | -30.4738
20.0952 | -35.2064 | -30.4808
20.1013 | -35.2069 | -30.4881
20.1074 | -35.2074 | -30.4952
20.1135 | -35.208 | -30.5021
20.1196 | -35.2085 | -30.5087
20.1257 | -35.2089 | -30.5157
20.1319 | -35.2095 | -30.5231
20.138 | -35.2099 | -30.5294
20.1441 | -35.2105 | -30.5365
20.1502 | -35.211 | -30.5432
20.1563 | -35.2116 | -30.5496
20.1624 | -35.212 | -30.5569
20.1685 | -35.2125 | -30.5636
20.1746 | -35.213 | -30.5709
20.1807 | -35.2135 | -30.5781
20.1868 | -35.2141 | -30.5853
20.1929 | -35.2146 | -30.5925
20.199 | -35.2152 | -30.5995
20.2051 | -35.2156 | -30.6062
20.2112 | -35.2161 | -30.6131
20.2173 | -35.2166 | -30.6199
20.2234 | -35.2171 | -30.6271
20.2295 | -35.2176 | -30.6341
20.2356 | -35.2182 | -30.641
20.2417 | -35.2186 | -30.647
20.2478 | -35.2191 | -30.654
20.2539 | -35.2196 | -30.6611
20.26 | -35.2201 | -30.6677
20.2661 | -35.2206 | -30.6749
20.2723 | -35.2211 | -30.6818
20.2784 | -35.2216 | -30.6883
20.2845 | -35.2221 | -30.6954
20.2906 | -35.2226 | -30.7018
20.2967 | -35.2231 | -30.7085
20.3028 | -35.2236 | -30.7153
20.3089 | -35.224 | -30.722
20.315 | -35.2245 | -30.7294
20.3211 | -35.225 | -30.736
20.3272 | -35.2255 | -30.7432
20.3333 | -35.2259 | -30.7496
20.3394 | -35.2264 | -30.7563
20.3455 | -35.2269 | -30.763
20.3516 | -35.2274 | -30.7698
20.3577 | -35.2279 | -30.7767
20.3638 | -35.2283 | -30.7837
20.3699 | -35.2289 | -30.7906
20.376 | -35.2294 | -30.7973
20.3821 | -35.2299 | -30.8044
20.3882 | -35.2303 | -30.8115
20.3943 | -35.2308 | -30.8182
20.4004 | -35.2313 | -30.8254
20.4065 | -35.2318 | -30.832
20.4126 | -35.2323 | -30.8391
20.4188 | -35.2328 | -30.846
20.4249 | -35.2333 | -30.8527
20.431 | -35.2338 | -30.8594
20.4371 | -35.2344 | -30.8661
20.4432 | -35.2348 | -30.8733
20.4493 | -35.2353 | -30.8802
20.4554 | -35.2359 | -30.8869
20.4615 | -35.2363 | -30.8939
20.4676 | -35.2369 | -30.901
20.4737 | -35.2374 | -30.9079
20.4798 | -35.2379 | -30.9146
20.4859 | -35.2384 | -30.9221
20.492 | -35.2389 | -30.9288
20.4981 | -35.2394 | -30.9356
20.5042 | -35.2399 | -30.9423
20.5103 | -35.2405 | -30.9479
20.5164 | -35.241 | -30.955
20.5225 | -35.2414 | -30.9615
20.5286 | -35.2419 | -30.9686
20.5347 | -35.2424 | -30.975
20.5408 | -35.2429 | -30.981
20.5469 | -35.2434 | -30.9877
20.553 | -35.2439 | -30.9946
20.5592 | -35.2444 | -31.0018
20.5653 | -35.2449 | -31.0084
20.5714 | -35.2454 | -31.0151
20.5775 | -35.246 | -31.0224
20.5836 | -35.2465 | -31.0299
20.5897 | -35.2472 | -31.0364
20.5958 | -35.2477 | -31.0429
20.6019 | -35.2481 | -31.05
20.608 | -35.2486 | -31.0571
20.6141 | -35.2491 | -31.0628
20.6202 | -35.2496 | -31.0696
20.6263 | -35.2501 | -31.0766
20.6324 | -35.2507 | -31.083
20.6385 | -35.2512 | -31.0902
20.6446 | -35.2518 | -31.0967
20.6507 | -35.2522 | -31.1035
20.6568 | -35.2528 | -31.11
20.6629 | -35.2534 | -31.1167
20.669 | -35.2539 | -31.1235
20.6751 | -35.2544 | -31.1306
20.6812 | -35.2547 | -31.1359
20.6873 | -35.2552 | -31.1425
20.6934 | -35.2558 | -31.1501
20.6995 | -35.2564 | -31.1573
20.7057 | -35.2569 | -31.1641
20.7118 | -35.2574 | -31.1707
20.7179 | -35.258 | -31.1777
20.724 | -35.2585 | -31.1849
20.7301 | -35.2589 | -31.1911
20.7362 | -35.2594 | -31.1976
20.7423 | -35.2599 | -31.2048
20.7484 | -35.2604 | -31.2111
20.7545 | -35.2609 | -31.217
20.7606 | -35.2613 | -31.2241
20.7667 | -35.2619 | -31.2315
20.7728 | -35.2624 | -31.2389
20.7789 | -35.2629 | -31.2457
20.785 | -35.2633 | -31.253
20.7911 | -35.2637 | -31.2601
20.7972 | -35.264 | -31.2674
20.8033 | -35.2646 | -31.2738
20.8094 | -35.2651 | -31.2803
20.8155 | -35.2655 | -31.2873
20.8216 | -35.266 | -31.2951
20.8277 | -35.2664 | -31.3014
20.8338 | -35.267 | -31.3084
20.8399 | -35.2675 | -31.3155
20.8461 | -35.2682 | -31.3218
20.8522 | -35.2687 | -31.3286
20.8583 | -35.2692 | -31.3339
20.8644 | -35.2697 | -31.3411
20.8705 | -35.2701 | -31.3479
20.8766 | -35.2706 | -31.355
20.8827 | -35.2711 | -31.3615
20.8888 | -35.2715 | -31.3684
20.8949 | -35.272 | -31.3758
20.901 | -35.2725 | -31.382
20.9071 | -35.273 | -31.3895
20.9132 | -35.2735 | -31.3971
20.9193 | -35.2739 | -31.4043
20.9254 | -35.2744 | -31.4114
20.9315 | -35.275 | -31.4185
20.9376 | -35.2755 | -31.4246
20.9437 | -35.276 | -31.4323
20.9498 | -35.2766 | -31.4382
20.9559 | -35.2771 | -31.4449
20.962 | -35.2777 | -31.4522
20.9681 | -35.2782 | -31.4582
20.9742 | -35.2786 | -31.466
20.9803 | -35.2792 | -31.4726
20.9864 | -35.2795 | -31.4789
20.9926 | -35.2801 | -31.4869
20.9987 | -35.2806 | -31.4957
21.0048 | -35.2812 | -31.5016
21.0109 | -35.2816 | -31.5083
21.017 | -35.282 | -31.515
21.0231 | -35.2825 | -31.5221
21.0292 | -35.283 | -31.5293
21.0353 | -35.2835 | -31.5356
21.0414 | -35.284 | -31.5435
21.0475 | -35.2846 | -31.5494
21.0536 | -35.2849 | -31.5573
21.0597 | -35.2854 | -31.5658
21.0658 | -35.2859 | -31.5716
21.0719 | -35.2863 | -31.5793
21.078 | -35.2867 | -31.5858
21.0841 | -35.2873 | -31.593
21.0902 | -35.2876 | -31.6004
21.0963 | -35.288 | -31.6066
21.1024 | -35.2885 | -31.6134
21.1085 | -35.289 | -31.6202
21.1146 | -35.2894 | -31.627
21.1207 | -35.2899 | -31.6336
21.1268 | -35.2904 | -31.6401
21.133 | -35.291 | -31.6467
21.1391 | -35.2914 | -31.6546
21.1452 | -35.2918 | -31.662
21.1513 | -35.2923 | -31.6689
21.1574 | -35.2928 | -31.6756
21.1635 | -35.2934 | -31.6822
21.1696 | -35.294 | -31.689
21.1757 | -35.2945 | -31.6966
21.1818 | -35.2948 | -31.7034
21.1879 | -35.2954 | -31.7088
21.194 | -35.296 | -31.7168
21.2001 | -35.2966 | -31.7238
21.2062 | -35.2969 | -31.7298
21.2123 | -35.2973 | -31.7362
21.2184 | -35.2979 | -31.743
21.2245 | -35.2983 | -31.7498
21.2306 | -35.2988 | -31.7567
21.2367 | -35.2994 | -31.7638
21.2428 | -35.2997 | -31.7701
21.2489 | -35.3002 | -31.7775
21.255 | -35.3007 | -31.7847
21.2611 | -35.3013 | -31.7921
21.2672 | -35.3019 | -31.7992
21.2733 | -35.3025 | -31.8064
21.2795 | -35.3031 | -31.8155
21.2856 | -35.3034 | -31.8224
21.2917 | -35.3039 | -31.8305
21.2978 | -35.3044 | -31.8355
21.3039 | -35.305 | -31.8418
21.31 | -35.3054 | -31.8502
21.3161 | -35.3058 | -31.8577
21.3222 | -35.3063 | -31.8638
21.3283 | -35.3067 | -31.8704
21.3344 | -35.3072 | -31.8766
21.3405 | -35.3078 | -31.8841
21.3466 | -35.3083 | -31.8914
21.3527 | -35.3089 | -31.8978
21.3588 | -35.3093 | -31.9053
21.3649 | -35.3096 | -31.9124
21.371 | -35.31 | -31.9193
21.3771 | -35.3105 | -31.9255
21.3832 | -35.3111 | -31.9327
21.3893 | -35.3116 | -31.9391
21.3954 | -35.3121 | -31.9457
21.4015 | -35.3127 | -31.952
21.4076 | -35.313 | -31.9589
21.4137 | -35.3135 | -31.9657
21.4199 | -35.314 | -31.9725
21.426 | -35.3146 | -31.9804
21.4321 | -35.3152 | -31.987
21.4382 | -35.3157 | -31.9943
21.4443 | -35.3163 | -32.0003
21.4504 | -35.3166 | -32.0068
21.4565 | -35.3172 | -32.014
21.4626 | -35.3175 | -32.0203
21.4687 | -35.3182 | -32.0275
21.4748 | -35.3187 | -32.0334
21.4809 | -35.3192 | -32.041
21.487 | -35.3197 | -32.047
21.4931 | -35.3203 | -32.0531
21.4992 | -35.3207 | -32.06
21.5053 | -35.3213 | -32.0674
21.5114 | -35.3219 | -32.0739
21.5175 | -35.3224 | -32.0807
21.5236 | -35.3229 | -32.0886
21.5297 | -35.3235 | -32.0948
21.5358 | -35.3238 | -32.1016
21.5419 | -35.3244 | -32.1083
21.548 | -35.3249 | -32.1151
21.5541 | -35.3253 | -32.1231
21.5602 | -35.326 | -32.1298
21.5664 | -35.3265 | -32.1376
21.5725 | -35.327 | -32.1446
21.5786 | -35.3274 | -32.1506
21.5847 | -35.3279 | -32.1589
21.5908 | -35.3284 | -32.1659
21.5969 | -35.3287 | -32.174
21.603 | -35.3291 | -32.1801
21.6091 | -35.3296 | -32.1864
21.6152 | -35.33 | -32.1928
21.6213 | -35.3307 | -32.2
21.6274 | -35.3312 | -32.2067
21.6335 | -35.3318 | -32.2131
21.6396 | -35.3323 | -32.2201
21.6457 | -35.3328 | -32.2264
21.6518 | -35.3332 | -32.233
21.6579 | -35.3338 | -32.2387
21.664 | -35.3343 | -32.2458
21.6701 | -35.3348 | -32.2516
21.6762 | -35.3352 | -32.2589
21.6823 | -35.3357 | -32.2668
21.6884 | -35.3364 | -32.2733
21.6945 | -35.337 | -32.2805
21.7006 | -35.3374 | -32.2873
21.7068 | -35.338 | -32.2938
21.7129 | -35.3385 | -32.3013
21.719 | -35.3389 | -32.3087
21.7251 | -35.3393 | -32.316
21.7312 | -35.3397 | -32.3225
21.7373 | -35.3403 | -32.3293
21.7434 | -35.3407 | -32.3356
21.7495 | -35.3413 | -32.3432
21.7556 | -35.3419 | -32.3502
21.7617 | -35.3424 | -32.3564
21.7678 | -35.3429 | -32.3639
21.7739 | -35.3433 | -32.3712
21.78 | -35.3437 | -32.3777
21.7861 | -35.3441 | -32.3847
21.7922 | -35.3446 | -32.392
21.7983 | -35.3451 | -32.3989
21.8044 | -35.3457 | -32.4069
21.8105 | -35.3461 | -32.4141
21.8166 | -35.3465 | -32.4214
21.8227 | -35.347 | -32.4287
21.8288 | -35.3474 | -32.4362
21.8349 | -35.3479 | -32.4421
21.841 | -35.3484 | -32.4485
21.8471 | -35.3491 | -32.4552
21.8533 | -35.3496 | -32.4624
21.8594 | -35.3501 | -32.4691
21.8655 | -35.3505 | -32.475
21.8716 | -35.3511 | -32.4811
21.8777 | -35.3516 | -32.4874
21.8838 | -35.3521 | -32.4942
21.8899 | -35.3527 | -32.5007
21.896 | -35.3532 | -32.5075
21.9021 | -35.3537 | -32.5148
21.9082 | -35.3542 | -32.5225
21.9143 | -35.3547 | -32.5289
21.9204 | -35.3553 | -32.5352
21.9265 | -35.3557 | -32.5419
21.9326 | -35.3563 | -32.5491
21.9387 | -35.3566 | -32.5559
21.9448 | -35.3573 | -32.5626
21.9509 | -35.358 | -32.5694
21.957 | -35.3584 | -32.5766
21.9631 | -35.3591 | -32.5838
21.9692 | -35.3596 | -32.5902
21.9753 | -35.36 | -32.5974
21.9814 | -35.3604 | -32.6045
21.9875 | -35.3609 | -32.6098
21.9937 | -35.3614 | -32.6177
21.9998 | -35.362 | -32.6241
22.0059 | -35.3626 | -32.6307
22.012 | -35.363 | -32.6383
22.0181 | -35.3634 | -32.6449
22.0242 | -35.3638 | -32.6511
22.0303 | -35.3643 | -32.658
22.0364 | -35.3648 | -32.6651
22.0425 | -35.3654 | -32.6718
22.0486 | -35.3659 | -32.6782
22.0547 | -35.3665 | -32.6841
22.0608 | -35.3671 | -32.6909
22.0669 | -35.3675 | -32.6972
22.073 | -35.3681 | -32.7036
22.0791 | -35.3687 | -32.7095
22.0852 | -35.3692 | -32.7171
22.0913 | -35.3697 | -32.7237
22.0974 | -35.3702 | -32.7306
22.1035 | -35.3706 | -32.7376
22.1096 | -35.3711 | -32.7447
22.1157 | -35.3715 | -32.7517
22.1218 | -35.372 | -32.7587
22.1279 | -35.3726 | -32.7653
22.134 | -35.373 | -32.7719
22.1402 | -35.3734 | -32.7791
22.1463 | -35.374 | -32.7857
22.1524 | -35.3745 | -32.7923
22.1585 | -35.3751 | -32.7988
22.1646 | -35.3756 | -32.8054
22.1707 | -35.3763 | -32.8123
22.1768 | -35.3767 | -32.8191
22.1829 | -35.3772 | -32.826
22.189 | -35.3778 | -32.8338
22.1951 | -35.3783 | -32.8407
22.2012 | -35.3788 | -32.8478
22.2073 | -35.3793 | -32.855
22.2134 | -35.3797 | -32.8616
22.2195 | -35.3802 | -32.8687
22.2256 | -35.3807 | -32.8755
22.2317 | -35.3813 | -32.8823
22.2378 | -35.3817 | -32.8888
22.2439 | -35.3823 | -32.8956
22.25 | -35.3827 | -32.9022
22.2561 | -35.3831 | -32.909
22.2622 | -35.3837 | -32.9164
22.2683 | -35.3842 | -32.9227
22.2744 | -35.3848 | -32.929
22.2806 | -35.3853 | -32.9353
22.2867 | -35.386 | -32.9421
22.2928 | -35.3865 | -32.9483
22.2989 | -35.387 | -32.956
22.305 | -35.3877 | -32.9635
22.3111 | -35.3882 | -32.9701
22.3172 | -35.3886 | -32.9771
22.3233 | -35.3892 | -32.9837
22.3294 | -35.3899 | -32.9896
22.3355 | -35.3904 | -32.9954
22.3416 | -35.391 | -33.0019
22.3477 | -35.3916 | -33.0081
22.3538 | -35.3921 | -33.0154
22.3599 | -35.3926 | -33.0226
22.366 | -35.3931 | -33.0298
22.3721 | -35.3936 | -33.0374
22.3782 | -35.3941 | -33.044
22.3843 | -35.3945 | -33.0507
22.3904 | -35.3951 | -33.0571
22.3965 | -35.3956 | -33.0642
22.4026 | -35.3961 | -33.0708
22.4087 | -35.3967 | -33.0777
22.4148 | -35.3972 | -33.0843
22.4209 | -35.3977 | -33.091
22.4271 | -35.3982 | -33.0978
22.4332 | -35.3987 | -33.1049
22.4393 | -35.3992 | -33.112
22.4454 | -35.3997 | -33.119
22.4515 | -35.4003 | -33.1251
22.4576 | -35.4009 | -33.1319
22.4637 | -35.4012 | -33.1383
22.4698 | -35.4018 | -33.1452
22.4759 | -35.4022 | -33.1512
22.482 | -35.4027 | -33.1581
22.4881 | -35.4031 | -33.1658
22.4942 | -35.4036 | -33.1727
22.5003 | -35.4042 | -33.1797
22.5064 | -35.4047 | -33.1864
22.5125 | -35.4052 | -33.1925
22.5186 | -35.4056 | -33.1995
22.5247 | -35.4062 | -33.2069
22.5308 | -35.4067 | -33.2131
22.5369 | -35.4071 | -33.2194
22.543 | -35.4077 | -33.2253
22.5491 | -35.4081 | -33.231
22.5552 | -35.4087 | -33.2384
22.5613 | -35.4092 | -33.2451
22.5675 | -35.4097 | -33.252
22.5736 | -35.4102 | -33.2586
22.5797 | -35.4107 | -33.2658
22.5858 | -35.4111 | -33.2713
22.5919 | -35.4116 | -33.2774
22.598 | -35.4121 | -33.2839
22.6041 | -35.4127 | -33.2899
22.6102 | -35.4132 | -33.2965
22.6163 | -35.4137 | -33.3024
22.6224 | -35.4141 | -33.3098
22.6285 | -35.4148 | -33.3161
22.6346 | -35.4153 | -33.3229
22.6407 | -35.4158 | -33.3295
22.6468 | -35.4164 | -33.3359
22.6529 | -35.4168 | -33.3423
22.659 | -35.4172 | -33.3488
22.6651 | -35.4178 | -33.355
22.6712 | -35.4183 | -33.3626
22.6773 | -35.4188 | -33.3691
22.6834 | -35.4193 | -33.3758
22.6895 | -35.4198 | -33.3816
22.6956 | -35.4203 | -33.3882
22.7017 | -35.4209 | -33.3944
22.7079 | -35.4215 | -33.4008
22.714 | -35.422 | -33.4072
22.7201 | -35.4226 | -33.414
22.7262 | -35.4231 | -33.4208
22.7323 | -35.4235 | -33.427
22.7384 | -35.4239 | -33.4343
22.7445 | -35.4243 | -33.441
22.7506 | -35.4248 | -33.4475
22.7567 | -35.4253 | -33.454
22.7628 | -35.4258 | -33.4611
22.7689 | -35.4263 | -33.4674
22.775 | -35.4269 | -33.4745
22.7811 | -35.4273 | -33.4818
22.7872 | -35.4278 | -33.4882
22.7933 | -35.4284 | -33.4942
22.7994 | -35.4289 | -33.5004
22.8055 | -35.4294 | -33.5069
22.8116 | -35.4299 | -33.514
22.8177 | -35.4304 | -33.5203
22.8238 | -35.4308 | -33.5271
22.8299 | -35.4312 | -33.5332
22.836 | -35.4317 | -33.5398
22.8421 | -35.4322 | -33.5461
22.8482 | -35.4327 | -33.5524
22.8544 | -35.4331 | -33.5589
22.8605 | -35.4337 | -33.5655
22.8666 | -35.4341 | -33.5721
22.8727 | -35.4346 | -33.5783
22.8788 | -35.4351 | -33.5848
22.8849 | -35.4356 | -33.5922
22.891 | -35.436 | -33.5996
22.8971 | -35.4366 | -33.6071
22.9032 | -35.437 | -33.6131
22.9093 | -35.4375 | -33.6192
22.9154 | -35.4381 | -33.6256
22.9215 | -35.4386 | -33.6318
22.9276 | -35.439 | -33.639
22.9337 | -35.4395 | -33.6455
22.9398 | -35.44 | -33.6525
22.9459 | -35.4405 | -33.6592
22.952 | -35.4409 | -33.6655
22.9581 | -35.4415 | -33.6721
22.9642 | -35.442 | -33.6776
22.9703 | -35.4425 | -33.6839
22.9764 | -35.4429 | -33.6898
22.9825 | -35.4435 | -33.6969
22.9886 | -35.444 | -33.7036
22.9948 | -35.4444 | -33.7106
23.0009 | -35.4447 | -33.7176
23.007 | -35.4451 | -33.7228
23.0131 | -35.4456 | -33.7286
23.0192 | -35.446 | -33.7352
23.0253 | -35.4465 | -33.7418
23.0314 | -35.447 | -33.7479
23.0375 | -35.4474 | -33.7541
23.0436 | -35.4479 | -33.7595
23.0497 | -35.4484 | -33.7653
23.0558 | -35.449 | -33.7716
23.0619 | -35.4495 | -33.7782
23.068 | -35.45 | -33.7843
23.0741 | -35.4504 | -33.7898
23.0802 | -35.451 | -33.796
23.0863 | -35.4514 | -33.8027
23.0924 | -35.4518 | -33.8089
23.0985 | -35.4523 | -33.8147
23.1046 | -35.4529 | -33.8209
23.1107 | -35.4534 | -33.8265
23.1168 | -35.4539 | -33.8327
23.1229 | -35.4544 | -33.8385
23.129 | -35.4549 | -33.8447
23.1351 | -35.4554 | -33.8509
23.1413 | -35.4559 | -33.8575
23.1474 | -35.4563 | -33.8633
23.1535 | -35.4567 | -33.8699
23.1596 | -35.4573 | -33.8763
23.1657 | -35.4577 | -33.8829
23.1718 | -35.4582 | -33.8887
23.1779 | -35.4587 | -33.8958
23.184 | -35.4591 | -33.9028
23.1901 | -35.4596 | -33.9097
23.1962 | -35.4601 | -33.9159
23.2023 | -35.4606 | -33.9221
23.2084 | -35.4611 | -33.929
23.2145 | -35.4616 | -33.9348
23.2206 | -35.4621 | -33.9413
23.2267 | -35.4626 | -33.9475
23.2328 | -35.4631 | -33.9542
23.2389 | -35.4635 | -33.9605
23.245 | -35.464 | -33.9668
23.2511 | -35.4644 | -33.9733
23.2572 | -35.4649 | -33.9795
23.2633 | -35.4654 | -33.986
23.2694 | -35.4658 | -33.9925
23.2755 | -35.4663 | -33.9992
23.2817 | -35.4668 | -34.0061
23.2878 | -35.4673 | -34.0118
23.2939 | -35.4678 | -34.0181
23.3 | -35.4682 | -34.0244
23.3061 | -35.4688 | -34.0307
23.3122 | -35.4693 | -34.0369
23.3183 | -35.4697 | -34.0427
23.3244 | -35.4702 | -34.0492
23.3305 | -35.4706 | -34.0555
23.3366 | -35.4711 | -34.0619
23.3427 | -35.4716 | -34.0681
23.3488 | -35.472 | -34.0742
23.3549 | -35.4725 | -34.0805
23.361 | -35.473 | -34.0866
23.3671 | -35.4735 | -34.0928
23.3732 | -35.474 | -34.0986
23.3793 | -35.4745 | -34.1045
23.3854 | -35.475 | -34.1108
23.3915 | -35.4754 | -34.116
23.3976 | -35.4759 | -34.1221
23.4037 | -35.4764 | -34.1284
23.4098 | -35.4769 | -34.135
23.4159 | -35.4773 | -34.1408
23.422 | -35.4778 | -34.147
23.4282 | -35.4783 | -34.1534
23.4343 | -35.4788 | -34.1596
23.4404 | -35.4793 | -34.1659
23.4465 | -35.4797 | -34.1724
23.4526 | -35.4802 | -34.1788
23.4587 | -35.4807 | -34.1847
23.4648 | -35.4812 | -34.1897
23.4709 | -35.4816 | -34.196
23.477 | -35.482 | -34.202
23.4831 | -35.4825 | -34.2083
23.4892 | -35.4829 | -34.2142
23.4953 | -35.4833 | -34.2203
23.5014 | -35.4838 | -34.2265
23.5075 | -35.4843 | -34.2327
23.5136 | -35.4848 | -34.2392
23.5197 | -35.4854 | -34.2454
23.5258 | -35.4858 | -34.2509
23.5319 | -35.4863 | -34.256
23.538 | -35.4868 | -34.2616
23.5441 | -35.4872 | -34.2676
23.5502 | -35.4877 | -34.2742
23.5563 | -35.4881 | -34.2803
23.5624 | -35.4885 | -34.2863
23.5686 | -35.489 | -34.2916
23.5747 | -35.4895 | -34.2979
23.5808 | -35.49 | -34.3038
23.5869 | -35.4905 | -34.3104
23.593 | -35.491 | -34.3162
23.5991 | -35.4915 | -34.3225
23.6052 | -35.4919 | -34.3276
23.6113 | -35.4924 | -34.3338
23.6174 | -35.4929 | -34.34
23.6235 | -35.4933 | -34.3459
23.6296 | -35.4938 | -34.3523
23.6357 | -35.4942 | -34.3585
23.6418 | -35.4947 | -34.364
23.6479 | -35.4952 | -34.3701
23.654 | -35.4957 | -34.3761
23.6601 | -35.4961 | -34.3818
23.6662 | -35.4965 | -34.3881
23.6723 | -35.4969 | -34.3935
23.6784 | -35.4974 | -34.4002
23.6845 | -35.4979 | -34.4061
23.6906 | -35.4984 | -34.4119
23.6967 | -35.4989 | -34.4178
23.7028 | -35.4994 | -34.4241
23.7089 | -35.4998 | -34.4291
23.7151 | -35.5003 | -34.4352
23.7212 | -35.5008 | -34.4414
23.7273 | -35.5012 | -34.4477
23.7334 | -35.5016 | -34.454
23.7395 | -35.5021 | -34.4598
23.7456 | -35.5026 | -34.4659
23.7517 | -35.503 | -34.4719
23.7578 | -35.5035 | -34.4775
23.7639 | -35.504 | -34.4835
23.77 | -35.5044 | -34.4894
23.7761 | -35.5049 | -34.4954
23.7822 | -35.5053 | -34.5001
23.7883 | -35.5058 | -34.5061
23.7944 | -35.5063 | -34.5123
23.8005 | -35.5067 | -34.5178
23.8066 | -35.5072 | -34.5237
23.8127 | -35.5077 | -34.5293
23.8188 | -35.5082 | -34.5354
23.8249 | -35.5086 | -34.5415
23.831 | -35.5091 | -34.5477
23.8371 | -35.5095 | -34.5538
23.8432 | -35.5098 | -34.5599
23.8493 | -35.5102 | -34.566
23.8555 | -35.5106 | -34.5717
23.8616 | -35.511 | -34.5783
23.8677 | -35.5116 | -34.5844
23.8738 | -35.512 | -34.5903
23.8799 | -35.5125 | -34.5967
23.886 | -35.5129 | -34.6031
23.8921 | -35.5134 | -34.6087
23.8982 | -35.5139 | -34.6145
23.9043 | -35.5145 | -34.6205
23.9104 | -35.5149 | -34.6261
23.9165 | -35.5154 | -34.6318
23.9226 | -35.516 | -34.6372
23.9287 | -35.5164 | -34.6431
23.9348 | -35.5169 | -34.6485
23.9409 | -35.5174 | -34.6541
23.947 | -35.5178 | -34.6596
23.9531 | -35.5183 | -34.6644
23.9592 | -35.5188 | -34.6706
23.9653 | -35.5193 | -34.6764
23.9714 | -35.5197 | -34.6821
23.9775 | -35.5202 | -34.6884
23.9836 | -35.5206 | -34.694
23.9897 | -35.5209 | -34.6993
23.9958 | -35.5214 | -34.7052
24.002 | -35.5218 | -34.7114
24.0081 | -35.5223 | -34.7176
24.0142 | -35.5228 | -34.7238
24.0203 | -35.5233 | -34.7296
24.0264 | -35.5237 | -34.7345
24.0325 | -35.5242 | -34.7401
24.0386 | -35.5247 | -34.7463
24.0447 | -35.5251 | -34.7516
24.0508 | -35.5255 | -34.7573
24.0569 | -35.5259 | -34.7629
24.063 | -35.5263 | -34.7685
24.0691 | -35.5268 | -34.7738
24.0752 | -35.5273 | -34.7798
24.0813 | -35.5278 | -34.7856
24.0874 | -35.5282 | -34.7916
24.0935 | -35.5287 | -34.7977
24.0996 | -35.5291 | -34.8028
24.1057 | -35.5295 | -34.8085
24.1118 | -35.5299 | -34.8147
24.1179 | -35.5304 | -34.8209
24.124 | -35.5307 | -34.827
24.1301 | -35.5312 | -34.8324
24.1362 | -35.5317 | -34.8381
24.1424 | -35.5321 | -34.8437
24.1485 | -35.5325 | -34.8497
24.1546 | -35.533 | -34.8551
24.1607 | -35.5335 | -34.8609
24.1668 | -35.5339 | -34.8656
24.1729 | -35.5343 | -34.8713
24.179 | -35.5348 | -34.8769
24.1851 | -35.5353 | -34.8824
24.1912 | -35.5357 | -34.8879
24.1973 | -35.5361 | -34.8937
24.2034 | -35.5366 | -34.8988
24.2095 | -35.537 | -34.9046
24.2156 | -35.5375 | -34.9105
24.2217 | -35.5379 | -34.9161
24.2278 | -35.5383 | -34.9223
24.2339 | -35.5387 | -34.9281
24.24 | -35.5391 | -34.9328
24.2461 | -35.5396 | -34.9385
24.2522 | -35.54 | -34.9438
24.2583 | -35.5404 | -34.9495
24.2644 | -35.5408 | -34.9553
24.2705 | -35.5413 | -34.96
24.2766 | -35.5417 | -34.9656
24.2827 | -35.5422 | -34.9709
24.2889 | -35.5427 | -34.9766
24.295 | -35.5431 | -34.982
24.3011 | -35.5435 | -34.9874
24.3072 | -35.5439 | -34.993
24.3133 | -35.5444 | -34.9988
24.3194 | -35.5448 | -35.0046
24.3255 | -35.5452 | -35.0104
24.3316 | -35.5457 | -35.0162
24.3377 | -35.5462 | -35.0214
24.3438 | -35.5465 | -35.0264
24.3499 | -35.5469 | -35.0316
24.356 | -35.5473 | -35.0367
24.3621 | -35.5476 | -35.0416
24.3682 | -35.5481 | -35.0475
24.3743 | -35.5485 | -35.0529
24.3804 | -35.5489 | -35.0576
24.3865 | -35.5494 | -35.0633
24.3926 | -35.5499 | -35.0688
24.3987 | -35.5503 | -35.074
24.4048 | -35.5507 | -35.0793
24.4109 | -35.5511 | -35.0845
24.417 | -35.5515 | -35.0902
24.4231 | -35.5519 | -35.0956
24.4293 | -35.5523 | -35.1009
24.4354 | -35.5527 | -35.1065
24.4415 | -35.5532 | -35.1121
24.4476 | -35.5535 | -35.1176
24.4537 | -35.554 | -35.1232
24.4598 | -35.5545 | -35.1288
24.4659 | -35.5548 | -35.1342
24.472 | -35.5553 | -35.1394
24.4781 | -35.5557 | -35.1448
24.4842 | -35.556 | -35.1502
24.4903 | -35.5564 | -35.1553
24.4964 | -35.5568 | -35.1604
24.5025 | -35.5572 | -35.1659
24.5086 | -35.5576 | -35.1712
24.5147 | -35.558 | -35.1767
24.5208 | -35.5585 | -35.1817
24.5269 | -35.5589 | -35.1871
24.533 | -35.5593 | -35.1927
24.5391 | -35.5598 | -35.1985
24.5452 | -35.5602 | -35.2041
24.5513 | -35.5606 | -35.2085
24.5574 | -35.5611 | -35.2135
24.5635 | -35.5616 | -35.2189
24.5696 | -35.562 | -35.2242
24.5758 | -35.5624 | -35.2298
24.5819 | -35.5629 | -35.2351
24.588 | -35.5633 | -35.2404
24.5941 | -35.5637 | -35.2457
24.6002 | -35.5642 | -35.2512
24.6063 | -35.5645 | -35.2561
24.6124 | -35.5649 | -35.2612
24.6185 | -35.5654 | -35.2666
24.6246 | -35.5658 | -35.2714
24.6307 | -35.5663 | -35.2767
24.6368 | -35.5667 | -35.2822
24.6429 | -35.5672 | -35.2877
24.649 | -35.5676 | -35.293
24.6551 | -35.568 | -35.2984
24.6612 | -35.5685 | -35.3027
24.6673 | -35.5689 | -35.3078
24.6734 | -35.5693 | -35.313
24.6795 | -35.5698 | -35.3182
24.6856 | -35.5702 | -35.3233
24.6917 | -35.5707 | -35.3289
24.6978 | -35.5711 | -35.3333
24.7039 | -35.5715 | -35.3388
24.71 | -35.572 | -35.3438
24.7162 | -35.5724 | -35.3493
24.7223 | -35.5728 | -35.3547
24.7284 | -35.5732 | -35.3596
24.7345 | -35.5736 | -35.3653
24.7406 | -35.5741 | -35.3708
24.7467 | -35.5744 | -35.3764
24.7528 | -35.575 | -35.3817
24.7589 | -35.5754 | -35.3873
24.765 | -35.5759 | -35.3917
24.7711 | -35.5763 | -35.3968
24.7772 | -35.5768 | -35.4021
24.7833 | -35.5771 | -35.4076
24.7894 | -35.5775 | -35.4128
24.7955 | -35.578 | -35.4181
24.8016 | -35.5783 | -35.4221
24.8077 | -35.5788 | -35.4273
24.8138 | -35.5792 | -35.4327
24.8199 | -35.5797 | -35.4379
24.826 | -35.5801 | -35.4432
24.8321 | -35.5805 | -35.4482
24.8382 | -35.5807 | -35.4523
24.8443 | -35.5812 | -35.4573
24.8504 | -35.5816 | -35.4624
24.8565 | -35.5821 | -35.4675
24.8627 | -35.5825 | -35.4729
24.8688 | -35.5829 | -35.4783
24.8749 | -35.5833 | -35.4837
24.881 | -35.5837 | -35.4892
24.8871 | -35.584 | -35.4949
24.8932 | -35.5845 | -35.4999
24.8993 | -35.5849 | -35.505
24.9054 | -35.5852 | -35.5096
24.9115 | -35.5856 | -35.5148
24.9176 | -35.586 | -35.5197
24.9237 | -35.5865 | -35.5254
24.9298 | -35.5868 | -35.5307
24.9359 | -35.5872 | -35.5363
24.942 | -35.5875 | -35.5414
24.9481 | -35.588 | -35.5464
24.9542 | -35.5884 | -35.5517
24.9603 | -35.5888 | -35.5567
24.9664 | -35.5892 | -35.562
24.9725 | -35.5896 | -35.5676
24.9786 | -35.5899 | -35.5724
24.9847 | -35.5903 | -35.5776
24.9908 | -35.5907 | -35.583
24.9969 | -35.5911 | -35.5884
25.0031 | -35.5915 | -35.5939
25.0092 | -35.5919 | -35.5987
25.0153 | -35.5923 | -35.6035
25.0214 | -35.5927 | -35.6089
25.0275 | -35.5931 | -35.6144
25.0336 | -35.5935 | -35.6198
25.0397 | -35.5939 | -35.625
25.0458 | -35.5943 | -35.6295
25.0519 | -35.5947 | -35.6346
25.058 | -35.5951 | -35.6397
25.0641 | -35.5955 | -35.645
25.0702 | -35.5959 | -35.6501
25.0763 | -35.5963 | -35.6552
25.0824 | -35.5967 | -35.6598
25.0885 | -35.5971 | -35.6646
25.0946 | -35.5975 | -35.67
25.1007 | -35.5978 | -35.6757
25.1068 | -35.5982 | -35.6804
25.1129 | -35.5986 | -35.6857
25.119 | -35.599 | -35.6911
25.1251 | -35.5994 | -35.6961
25.1312 | -35.5998 | -35.7015
25.1373 | -35.6002 | -35.7064
25.1435 | -35.6006 | -35.7112
25.1496 | -35.6011 | -35.7162
25.1557 | -35.6014 | -35.7213
25.1618 | -35.6018 | -35.7267
25.1679 | -35.6021 | -35.732
25.174 | -35.6024 | -35.737
25.1801 | -35.6029 | -35.7425
25.1862 | -35.6033 | -35.7473
25.1923 | -35.6036 | -35.7526
25.1984 | -35.604 | -35.758
25.2045 | -35.6044 | -35.763
25.2106 | -35.6048 | -35.7681
25.2167 | -35.6052 | -35.7733
25.2228 | -35.6055 | -35.7781
25.2289 | -35.6059 | -35.7828
25.235 | -35.6063 | -35.7878
25.2411 | -35.6066 | -35.7929
25.2472 | -35.6069 | -35.798
25.2533 | -35.6073 | -35.8033
25.2594 | -35.6076 | -35.8086
25.2655 | -35.6079 | -35.8135
25.2716 | -35.6083 | -35.8186
25.2777 | -35.6087 | -35.8238
25.2838 | -35.6091 | -35.8288
25.29 | -35.6095 | -35.8341
25.2961 | -35.6098 | -35.8382
25.3022 | -35.6101 | -35.8432
25.3083 | -35.6105 | -35.848
25.3144 | -35.6109 | -35.853
25.3205 | -35.6113 | -35.8581
25.3266 | -35.6117 | -35.8626
25.3327 | -35.6121 | -35.8674
25.3388 | -35.6124 | -35.8722
25.3449 | -35.6128 | -35.8768
25.351 | -35.6131 | -35.8817
25.3571 | -35.6135 | -35.8868
25.3632 | -35.6139 | -35.8921
25.3693 | -35.6142 | -35.897
25.3754 | -35.6146 | -35.902
25.3815 | -35.6149 | -35.907
25.3876 | -35.6153 | -35.9124
25.3937 | -35.6157 | -35.9176
25.3998 | -35.616 | -35.9226
25.4059 | -35.6165 | -35.9278
25.412 | -35.6169 | -35.9329
25.4181 | -35.6173 | -35.9376
25.4242 | -35.6177 | -35.9429
25.4304 | -35.618 | -35.948
25.4365 | -35.6184 | -35.9526
25.4426 | -35.6187 | -35.9577
25.4487 | -35.6191 | -35.9629
25.4548 | -35.6194 | -35.9676
25.4609 | -35.6198 | -35.9725
25.467 | -35.6201 | -35.9768
25.4731 | -35.6205 | -35.9819
25.4792 | -35.6208 | -35.987
25.4853 | -35.6211 | -35.9918
25.4914 | -35.6215 | -35.9969
25.4975 | -35.6218 | -36.0017
25.5036 | -35.6221 | -36.0065
25.5097 | -35.6226 | -36.0113
25.5158 | -35.6229 | -36.0166
25.5219 | -35.6233 | -36.0218
25.528 | -35.6237 | -36.0268
25.5341 | -35.6241 | -36.032
25.5402 | -35.6244 | -36.0368
25.5463 | -35.6249 | -36.0417
25.5524 | -35.6253 | -36.0465
25.5585 | -35.6256 | -36.0517
25.5646 | -35.6261 | -36.0565
25.5707 | -35.6265 | -36.0613
25.5769 | -35.6267 | -36.0661
25.583 | -35.6271 | -36.0713
25.5891 | -35.6274 | -36.076
25.5952 | -35.6278 | -36.0809
25.6013 | -35.6282 | -36.0854
25.6074 | -35.6285 | -36.0903
25.6135 | -35.6289 | -36.0956
25.6196 | -35.6292 | -36.1007
25.6257 | -35.6296 | -36.1053
25.6318 | -35.6299 | -36.1102
25.6379 | -35.6303 | -36.1152
25.644 | -35.6307 | -36.1205
25.6501 | -35.6311 | -36.125
25.6562 | -35.6314 | -36.1296
25.6623 | -35.6318 | -36.1349
25.6684 | -35.6321 | -36.1399
25.6745 | -35.6325 | -36.1447
25.6806 | -35.6329 | -36.1496
25.6867 | -35.6333 | -36.1548
25.6928 | -35.6336 | -36.1599
25.6989 | -35.634 | -36.1654
25.705 | -35.6344 | -36.1705
25.7111 | -35.6347 | -36.1753
25.7173 | -35.635 | -36.1801
25.7234 | -35.6353 | -36.185
25.7295 | -35.6357 | -36.1902
25.7356 | -35.636 | -36.1953
25.7417 | -35.6364 | -36.2003
25.7478 | -35.6367 | -36.2053
25.7539 | -35.6371 | -36.2107
25.76 | -35.6374 | -36.2158
25.7661 | -35.6378 | -36.2203
25.7722 | -35.6382 | -36.2252
25.7783 | -35.6386 | -36.2302
25.7844 | -35.6389 | -36.2348
25.7905 | -35.6392 | -36.2396
25.7966 | -35.6395 | -36.2447
25.8027 | -35.6398 | -36.2495
25.8088 | -35.6401 | -36.2545
25.8149 | -35.6405 | -36.2596
25.821 | -35.6408 | -36.2651
25.8271 | -35.6412 | -36.2697
25.8332 | -35.6415 | -36.2748
25.8393 | -35.6418 | -36.2797
25.8454 | -35.6422 | -36.285
25.8515 | -35.6425 | -36.29
25.8576 | -35.6428 | -36.2949
25.8638 | -35.6431 | -36.2998
25.8699 | -35.6434 | -36.3049
25.876 | -35.6437 | -36.3103
25.8821 | -35.6441 | -36.3153
25.8882 | -35.6444 | -36.3206
25.8943 | -35.6448 | -36.3256
25.9004 | -35.6451 | -36.3303
25.9065 | -35.6454 | -36.3354
25.9126 | -35.6457 | -36.34
25.9187 | -35.646 | -36.3452
25.9248 | -35.6463 | -36.3501
25.9309 | -35.6466 | -36.3552
25.937 | -35.647 | -36.3599
25.9431 | -35.6474 | -36.365
25.9492 | -35.6477 | -36.3701
25.9553 | -35.648 | -36.3752
25.9614 | -35.6483 | -36.3799
25.9675 | -35.6487 | -36.385
25.9736 | -35.649 | -36.39
25.9797 | -35.6493 | -36.3951
25.9858 | -35.6497 | -36.4001
25.9919 | -35.6501 | -36.4051
25.998 | -35.6503 | -36.4096
26.0042 | -35.6507 | -36.4145
26.0103 | -35.651 | -36.4191
26.0164 | -35.6513 | -36.4239
26.0225 | -35.6516 | -36.429
26.0286 | -35.652 | -36.4337
26.0347 | -35.6523 | -36.4385
26.0408 | -35.6526 | -36.4436
26.0469 | -35.6529 | -36.4483
26.053 | -35.6533 | -36.4532
26.0591 | -35.6536 | -36.4581
26.0652 | -35.6538 | -36.463
26.0713 | -35.6542 | -36.468
26.0774 | -35.6545 | -36.4727
26.0835 | -35.6548 | -36.478
26.0896 | -35.6551 | -36.4829
26.0957 | -35.6555 | -36.4875
26.1018 | -35.6559 | -36.4924
26.1079 | -35.6563 | -36.4974
26.114 | -35.6566 | -36.5024
26.1201 | -35.657 | -36.5073
26.1262 | -35.6573 | -36.5121
26.1323 | -35.6577 | -36.5169
26.1384 | -35.6579 | -36.5212
26.1445 | -35.6582 | -36.5265
26.1507 | -35.6585 | -36.5312
26.1568 | -35.6589 | -36.5359
26.1629 | -35.6592 | -36.5408
26.169 | -35.6595 | -36.5456
26.1751 | -35.6599 | -36.5507
26.1812 | -35.6603 | -36.5557
26.1873 | -35.6606 | -36.5608
26.1934 | -35.6609 | -36.5653
26.1995 | -35.6613 | -36.5706
26.2056 | -35.6616 | -36.5762
26.2117 | -35.6619 | -36.581
26.2178 | -35.6622 | -36.5859
26.2239 | -35.6625 | -36.5907
26.23 | -35.6628 | -36.5957
26.2361 | -35.6631 | -36.6005
26.2422 | -35.6634 | -36.6053
26.2483 | -35.6638 | -36.6099
26.2544 | -35.6641 | -36.6149
26.2605 | -35.6644 | -36.6197
26.2666 | -35.6647 | -36.6245
26.2727 | -35.6651 | -36.6294
26.2788 | -35.6654 | -36.6348
26.2849 | -35.6658 | -36.6401
26.2911 | -35.6661 | -36.645
26.2972 | -35.6665 | -36.6502
26.3033 | -35.6668 | -36.6551
26.3094 | -35.6671 | -36.6601
26.3155 | -35.6674 | -36.6649
26.3216 | -35.6678 | -36.6698
26.3277 | -35.6681 | -36.6748
26.3338 | -35.6683 | -36.6799
26.3399 | -35.6686 | -36.685
26.346 | -35.669 | -36.69
26.3521 | -35.6693 | -36.6951
26.3582 | -35.6696 | -36.7001
26.3643 | -35.6699 | -36.7049
26.3704 | -35.6703 | -36.7104
26.3765 | -35.6705 | -36.7157
26.3826 | -35.6709 | -36.7207
26.3887 | -35.6712 | -36.7258
26.3948 | -35.6715 | -36.7307
26.4009 | -35.6718 | -36.7358
26.407 | -35.6721 | -36.7409
26.4131 | -35.6724 | -36.746
26.4192 | -35.6727 | -36.7508
26.4253 | -35.6731 | -36.7558
26.4314 | -35.6734 | -36.7608
26.4376 | -35.6738 | -36.7659
26.4437 | -35.6741 | -36.7709
26.4498 | -35.6745 | -36.776
26.4559 | -35.6748 | -36.7812
26.462 | -35.6751 | -36.7863
26.4681 | -35.6754 | -36.7912
26.4742 | -35.6757 | -36.7961
26.4803 | -35.676 | -36.8013
26.4864 | -35.6763 | -36.8064
26.4925 | -35.6766 | -36.8113
26.4986 | -35.6769 | -36.8164
26.5047 | -35.6773 | -36.8216
26.5108 | -35.6776 | -36.8269
26.5169 | -35.6779 | -36.8316
26.523 | -35.6782 | -36.8369
26.5291 | -35.6785 | -36.8417
26.5352 | -35.6788 | -36.8467
26.5413 | -35.679 | -36.8516
26.5474 | -35.6793 | -36.8569
26.5535 | -35.6796 | -36.8621
26.5596 | -35.6799 | -36.8668
26.5657 | -35.6802 | -36.8716
26.5718 | -35.6805 | -36.8764
26.578 | -35.6808 | -36.881
26.5841 | -35.6811 | -36.8857
26.5902 | -35.6814 | -36.8905
26.5963 | -35.6817 | -36.8957
26.6024 | -35.6819 | -36.9006
26.6085 | -35.6822 | -36.9058
26.6146 | -35.6825 | -36.911
26.6207 | -35.6828 | -36.9161
26.6268 | -35.6832 | -36.9208
26.6329 | -35.6835 | -36.9258
26.639 | -35.6838 | -36.9307
26.6451 | -35.6841 | -36.9355
26.6512 | -35.6844 | -36.9406
26.6573 | -35.6848 | -36.9455
26.6634 | -35.6851 | -36.9507
26.6695 | -35.6854 | -36.9558
26.6756 | -35.6857 | -36.9611
26.6817 | -35.686 | -36.9664
26.6878 | -35.6863 | -36.9714
26.6939 | -35.6866 | -36.9766
26.7 | -35.6869 | -36.9818
26.7061 | -35.6872 | -36.9867
26.7122 | -35.6875 | -36.9918
26.7183 | -35.6878 | -36.9967
26.7245 | -35.6882 | -37.0019
26.7306 | -35.6885 | -37.0068
26.7367 | -35.6888 | -37.012
26.7428 | -35.6891 | -37.0169
26.7489 | -35.6894 | -37.022
26.755 | -35.6897 | -37.027
26.7611 | -35.69 | -37.0323
26.7672 | -35.6904 | -37.0373
26.7733 | -35.6906 | -37.0424
26.7794 | -35.691 | -37.0474
26.7855 | -35.6913 | -37.0525
26.7916 | -35.6915 | -37.057
26.7977 | -35.6919 | -37.0619
26.8038 | -35.6922 | -37.0672
26.8099 | -35.6925 | -37.072
26.816 | -35.6928 | -37.077
26.8221 | -35.693 | -37.082
26.8282 | -35.6934 | -37.0872
26.8343 | -35.6937 | -37.0922
26.8404 | -35.694 | -37.0974
26.8465 | -35.6943 | -37.1028
26.8526 | -35.6946 | -37.1079
26.8587 | -35.6949 | -37.113
26.8649 | -35.6952 | -37.1177
26.871 | -35.6955 | -37.1227
26.8771 | -35.6957 | -37.128
26.8832 | -35.6961 | -37.1329
26.8893 | -35.6964 | -37.138
26.8954 | -35.6967 | -37.1432
26.9015 | -35.697 | -37.148
26.9076 | -35.6973 | -37.153
26.9137 | -35.6976 | -37.1581
26.9198 | -35.6979 | -37.1632
26.9259 | -35.6982 | -37.1683
26.932 | -35.6984 | -37.1734
26.9381 | -35.6987 | -37.1784
26.9442 | -35.699 | -37.1839
26.9503 | -35.6994 | -37.1888
26.9564 | -35.6996 | -37.194
26.9625 | -35.7 | -37.199
26.9686 | -35.7002 | -37.2036
26.9747 | -35.7005 | -37.2091
26.9808 | -35.7008 | -37.2138
26.9869 | -35.701 | -37.2189
26.993 | -35.7013 | -37.2239
26.9991 | -35.7016 | -37.2286
27.0052 | -35.7019 | -37.2338
27.0114 | -35.7021 | -37.2392
27.0175 | -35.7024 | -37.2442
27.0236 | -35.7028 | -37.2493
27.0297 | -35.703 | -37.2545
27.0358 | -35.7033 | -37.2594
27.0419 | -35.7036 | -37.2645
27.048 | -35.7038 | -37.2697
27.0541 | -35.7041 | -37.2751
27.0602 | -35.7044 | -37.28
27.0663 | -35.7047 | -37.2853
27.0724 | -35.7049 | -37.2906
27.0785 | -35.7052 | -37.2961
27.0846 | -35.7055 | -37.3013
27.0907 | -35.7058 | -37.3064
27.0968 | -35.7061 | -37.3117
27.1029 | -35.7064 | -37.3169
27.109 | -35.7067 | -37.3218
27.1151 | -35.707 | -37.3267
27.1212 | -35.7072 | -37.332
27.1273 | -35.7075 | -37.3373
27.1334 | -35.7078 | -37.3424
27.1395 | -35.7081 | -37.3476
27.1456 | -35.7084 | -37.3528
27.1518 | -35.7086 | -37.3579
27.1579 | -35.7089 | -37.3629
27.164 | -35.7092 | -37.3678
27.1701 | -35.7095 | -37.3728
27.1762 | -35.7098 | -37.3779
27.1823 | -35.71 | -37.3827
27.1884 | -35.7103 | -37.3881
27.1945 | -35.7106 | -37.3934
27.2006 | -35.7109 | -37.3986
27.2067 | -35.7112 | -37.4035
27.2128 | -35.7115 | -37.4088
27.2189 | -35.7118 | -37.4139
27.225 | -35.7121 | -37.4191
27.2311 | -35.7124 | -37.4244
27.2372 | -35.7127 | -37.4295
27.2433 | -35.713 | -37.4343
27.2494 | -35.7132 | -37.4395
27.2555 | -35.7135 | -37.4444
27.2616 | -35.7138 | -37.4491
27.2677 | -35.714 | -37.4542
27.2738 | -35.7143 | -37.4593
27.2799 | -35.7145 | -37.4643
27.286 | -35.7148 | -37.4692
27.2921 | -35.7151 | -37.4742
27.2983 | -35.7154 | -37.4795
27.3044 | -35.7157 | -37.4848
27.3105 | -35.7159 | -37.4897
27.3166 | -35.7162 | -37.4945
27.3227 | -35.7165 | -37.4996
27.3288 | -35.7168 | -37.5048
27.3349 | -35.717 | -37.5096
27.341 | -35.7173 | -37.5147
27.3471 | -35.7176 | -37.5199
27.3532 | -35.7179 | -37.5251
27.3593 | -35.7181 | -37.5302
27.3654 | -35.7184 | -37.5353
27.3715 | -35.7186 | -37.5402
27.3776 | -35.7189 | -37.5455
27.3837 | -35.7192 | -37.5506
27.3898 | -35.7194 | -37.5557
27.3959 | -35.7197 | -37.5611
27.402 | -35.72 | -37.5662
27.4081 | -35.7203 | -37.5713
27.4142 | -35.7206 | -37.5763
27.4203 | -35.7208 | -37.5815
27.4264 | -35.7211 | -37.5867
27.4325 | -35.7214 | -37.5918
27.4387 | -35.7216 | -37.597
27.4448 | -35.7218 | -37.6018
27.4509 | -35.7221 | -37.6071
27.457 | -35.7224 | -37.6123
27.4631 | -35.7226 | -37.617
27.4692 | -35.7229 | -37.6221
27.4753 | -35.7231 | -37.6271
27.4814 | -35.7234 | -37.6323
27.4875 | -35.7237 | -37.6375
27.4936 | -35.7239 | -37.6424
27.4997 | -35.7242 | -37.6476
27.5058 | -35.7244 | -37.6528
27.5119 | -35.7247 | -37.6582
27.518 | -35.725 | -37.6632
27.5241 | -35.7253 | -37.6684
27.5302 | -35.7256 | -37.6738
27.5363 | -35.7258 | -37.6791
27.5424 | -35.7261 | -37.6844
27.5485 | -35.7264 | -37.6896
27.5546 | -35.7267 | -37.6944
27.5607 | -35.727 | -37.6998
27.5668 | -35.7272 | -37.7053
27.5729 | -35.7275 | -37.7105
27.5791 | -35.7278 | -37.7158
27.5852 | -35.7281 | -37.7211
27.5913 | -35.7283 | -37.7259
27.5974 | -35.7286 | -37.7309
27.6035 | -35.7289 | -37.7363
27.6096 | -35.7291 | -37.7413
27.6157 | -35.7293 | -37.7461
27.6218 | -35.7297 | -37.7513
27.6279 | -35.7299 | -37.7562
27.634 | -35.7302 | -37.7612
27.6401 | -35.7305 | -37.7663
27.6462 | -35.7307 | -37.7714
27.6523 | -35.7309 | -37.7766
27.6584 | -35.7313 | -37.7818
27.6645 | -35.7315 | -37.787
27.6706 | -35.7318 | -37.7921
27.6767 | -35.732 | -37.7977
27.6828 | -35.7323 | -37.8027
27.6889 | -35.7326 | -37.8077
27.695 | -35.7329 | -37.8134
27.7011 | -35.7331 | -37.8186
27.7072 | -35.7334 | -37.8236
27.7133 | -35.7337 | -37.8293
27.7194 | -35.734 | -37.8343
27.7256 | -35.7342 | -37.8396
27.7317 | -35.7345 | -37.8448
27.7378 | -35.7347 | -37.85
27.7439 | -35.735 | -37.8553
27.75 | -35.7353 | -37.8606
27.7561 | -35.7356 | -37.866
27.7622 | -35.7359 | -37.8713
27.7683 | -35.7362 | -37.8766
27.7744 | -35.7365 | -37.8821
27.7805 | -35.7367 | -37.8868
27.7866 | -35.737 | -37.892
27.7927 | -35.7372 | -37.8973
27.7988 | -35.7374 | -37.9022
27.8049 | -35.7377 | -37.9073
27.811 | -35.738 | -37.9121
27.8171 | -35.7382 | -37.9172
27.8232 | -35.7385 | -37.9224
27.8293 | -35.7388 | -37.9277
27.8354 | -35.7391 | -37.9328
27.8415 | -35.7393 | -37.9376
27.8476 | -35.7396 | -37.9425
27.8537 | -35.7398 | -37.9474
27.8598 | -35.74 | -37.9527
27.866 | -35.7403 | -37.9583
27.8721 | -35.7405 | -37.9634
27.8782 | -35.7408 | -37.9686
27.8843 | -35.741 | -37.9738
27.8904 | -35.7413 | -37.9787
27.8965 | -35.7416 | -37.9835
27.9026 | -35.7419 | -37.9883
27.9087 | -35.7422 | -37.9937
27.9148 | -35.7425 | -37.9994
27.9209 | -35.7427 | -38.0044
27.927 | -35.7429 | -38.0088
27.9331 | -35.7431 | -38.0141
27.9392 | -35.7434 | -38.0194
27.9453 | -35.7438 | -38.0246
27.9514 | -35.744 | -38.0299
27.9575 | -35.7442 | -38.0351
27.9636 | -35.7444 | -38.0404
27.9697 | -35.7447 | -38.0454
27.9758 | -35.745 | -38.051
27.9819 | -35.7451 | -38.0562
27.988 | -35.7453 | -38.0609
27.9941 | -35.7455 | -38.0658
28.0002 | -35.7458 | -38.071
28.0063 | -35.7461 | -38.0761
28.0125 | -35.7464 | -38.0812
28.0186 | -35.7466 | -38.0867
28.0247 | -35.7469 | -38.0922
28.0308 | -35.7471 | -38.0968
28.0369 | -35.7473 | -38.1018
28.043 | -35.7476 | -38.107
28.0491 | -35.748 | -38.1119
28.0552 | -35.7482 | -38.1168
28.0613 | -35.7485 | -38.1216
28.0674 | -35.7488 | -38.1271
28.0735 | -35.749 | -38.132

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4743 | -42.7417
21.4908 | -35.5836 | -35.1728
16.4948 | -34.9512 | -27.81
12.6602 | -34.5831 | -21.53
9.7171 | -34.3608 | -16.6172
7.4582 | -34.2099 | -12.6302
5.7244 | -34.106 | -9.4118
4.3936 | -34.0381 | -6.7956
3.3722 | -34.0084 | -4.7432
2.5883 | -33.9919 | -3.1837
1.9866 | -33.9942 | -1.9434
1.5248 | -33.9823 | -0.94274
1.1703 | -34.0114 | -0.29594
1 | -34.0038 | 0.028264
0.89822 | -33.9837 | 0.30286
0.68941 | -33.9686 | 0.88213
0.52919 | -33.9788 | 1.3348
0.40612 | -33.9832 | 1.9407
0.31171 | -33.9757 | 2.3925
0.23928 | -33.9708 | 3.1159
0.18358 | -33.9668 | 3.5388
0.14091 | -33.9887 | 4.4085
0.10819 | -34.0032 | 5.5636
0.083017 | -34.0259 | 7.0305
0.063705 | -34.0846 | 8.8769
0.048923 | -34.1451 | 11.331
0.037479 | -34.2412 | 14.5059
0.028801 | -34.4204 | 18.557
0.022078 | -34.7435 | 24.0392
0.017023 | -35.2342 | 30.4447
0.013018 | -36.0396 | 38.1093
0.010014 | -37.141 | 46.0178

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003865 | -6.8129
0.007729 | -26.088
0.011594 | -34.5838
0.015459 | -39.7023
0.019324 | -42.7116
0.023188 | -44.3427
0.027053 | -46.8049
0.030918 | -49.0206
0.034783 | -50.3642
0.038647 | -50.183
0.042512 | -51.7379
0.046377 | -53.1933
0.050242 | -53.5028
0.054106 | -54.2401
0.057971 | -54.741
0.061836 | -54.3784
0.0657 | -56.3875
0.069565 | -58.9695
0.07343 | -59.5806
0.077295 | -59.0901
0.081159 | -58.9609
0.085024 | -59.9684
0.088889 | -59.9446
0.092754 | -60.124
0.096618 | -61.3745
0.10048 | -61.4289
0.10435 | -61.6474
0.10821 | -62.4818
0.11208 | -62.2627
0.11594 | -63.4202
0.11981 | -63.1225
0.12367 | -64.0878
0.12754 | -65.1657
0.1314 | -64.5858
0.13527 | -64.1772
0.13913 | -64.4579
0.143 | -65.8336
0.14686 | -67.1587
0.15072 | -66.289
0.15459 | -65.8318
0.15845 | -66.7047
0.16232 | -67.9392
0.16618 | -67.228
0.17005 | -66.9995
0.17391 | -67.3236
0.17778 | -67.7285
0.18164 | -69.1144
0.18551 | -69.1353
0.18937 | -68.9789
0.19324 | -68.9958
0.1971 | -69.1081
0.20097 | -68.7438
0.20483 | -69.5243
0.2087 | -69.7671
0.21256 | -69.5362
0.21643 | -69.6612
0.22029 | -70.4929
0.22415 | -69.8263
0.22802 | -69.7139
0.23188 | -70.6902
0.23575 | -71.1993
0.23961 | -72.02
0.24348 | -72.2983
0.24734 | -72.1698
0.25121 | -73.2182
0.25507 | -72.5345
0.25894 | -72.6086
0.2628 | -72.1761
0.26667 | -71.7073
0.27053 | -71.1818
0.2744 | -71.8377
0.27826 | -73.466
0.28213 | -73.2196
0.28599 | -72.9592
0.28985 | -72.6246
0.29372 | -72.9196
0.29758 | -74.6092
0.30145 | -74.9979
0.30531 | -74.2361
0.30918 | -73.155
0.31304 | -73.9671
0.31691 | -73.9793
0.32077 | -73.9765
0.32464 | -74.6081
0.3285 | -74.7505
0.33237 | -74.6199
0.33623 | -74.4318
0.3401 | -76.3151
0.34396 | -76.4552
0.34783 | -75.1226
0.35169 | -75.2275
0.35556 | -76.1521
0.35942 | -76.5356
0.36329 | -77.5274
0.36715 | -75.9903
0.37101 | -75.954
0.37488 | -75.9254
0.37874 | -76.1879
0.38261 | -75.5798
0.38647 | -76.008
0.39034 | -77.2516
0.3942 | -76.7658
0.39807 | -77.1878
0.40193 | -76.9444
0.4058 | -76.3066
0.40966 | -76.8832
0.41353 | -76.2761
0.41739 | -76.472
0.42126 | -77.1135
0.42512 | -77.0894
0.42899 | -77.1025
0.43285 | -77.8778
0.43672 | -78.2997
0.44058 | -77.2333
0.44444 | -76.9697
0.44831 | -78.3369
0.45217 | -79.4718
0.45604 | -79.1651
0.4599 | -78.8039
0.46377 | -78.3475
0.46763 | -78.6844
0.4715 | -78.0074
0.47536 | -78.0191
0.47923 | -78.1663
0.48309 | -78.6452
0.48696 | -79.0503
0.49082 | -78.7432
0.49469 | -79.0421
0.49855 | -77.7664
0.50241 | -78.0966
0.50628 | -78.8839
0.51014 | -78.5836
0.51401 | -78.6906
0.51787 | -79.2146
0.52174 | -78.4846
0.5256 | -78.6622
0.52947 | -79.6766
0.53333 | -80.4199
0.5372 | -80.2615
0.54106 | -79.8048
0.54493 | -80.6337
0.54879 | -80.4165
0.55266 | -80.9357
0.55652 | -80.6843
0.56039 | -79.7891
0.56425 | -79.2101
0.56812 | -79.1568
0.57198 | -79.4945
0.57585 | -80.1971
0.57971 | -80.9682
0.58357 | -81.1354
0.58744 | -80.9533
0.5913 | -80.8425
0.59517 | -80.8073
0.59903 | -80.5095
0.6029 | -80.377
0.60676 | -80.4621
0.61063 | -80.8129
0.61449 | -81.5338
0.61836 | -81.8771
0.62222 | -81.3488
0.62609 | -80.4734
0.62995 | -80.1674
0.63382 | -80.5452
0.63768 | -81.0053
0.64155 | -81.3585
0.64541 | -81.3853
0.64928 | -81.3859
0.65314 | -81.2816
0.657 | -81.1935
0.66087 | -81.2133
0.66473 | -81.4284
0.6686 | -81.6412
0.67246 | -81.5494
0.67633 | -81.2793
0.68019 | -81.042
0.68406 | -81.1697
0.68792 | -81.6641
0.69179 | -82.093
0.69565 | -82.0885
0.69952 | -81.7514
0.70338 | -81.5841
0.70725 | -81.9346
0.71111 | -82.1995
0.71498 | -82.411
0.71884 | -82.7736
0.72271 | -82.8594
0.72657 | -82.6577
0.73043 | -82.1991
0.7343 | -81.7227
0.73816 | -81.8093
0.74203 | -82.2671
0.74589 | -82.4677
0.74976 | -82.5141
0.75362 | -82.4084
0.75749 | -82.5494
0.76135 | -82.7689
0.76522 | -82.7148
0.76908 | -82.6697
0.77295 | -82.6583
0.77681 | -82.6322
0.78068 | -82.7266
0.78454 | -82.92
0.78841 | -82.6836
0.79227 | -82.2998
0.79614 | -82.4929
0.8 | -82.9082
0.80387 | -83.6207
0.80773 | -83.7687
0.81159 | -83.319
0.81546 | -83.3846
0.81932 | -84.0206
0.82319 | -84.5786
0.82705 | -84.4457
0.83092 | -84.2343
0.83478 | -84.6402
0.83865 | -84.7228
0.84251 | -84.2816
0.84638 | -83.9008
0.85024 | -83.6432
0.85411 | -83.3487
0.85797 | -83.4965
0.86184 | -83.8317
0.8657 | -84.11
0.86957 | -84.5827
0.87343 | -84.9745
0.8773 | -84.9156
0.88116 | -84.5495
0.88502 | -84.2775
0.88889 | -84.3936
0.89275 | -84.6481
0.89662 | -84.4666
0.90048 | -84.3666
0.90435 | -84.6881
0.90821 | -85.0555
0.91208 | -85.1359
0.91594 | -85.28
0.91981 | -85.5397
0.92367 | -85.443
0.92754 | -85.1289
0.9314 | -84.7649
0.93527 | -84.5613
0.93913 | -84.2472
0.943 | -84.0219
0.94686 | -84.113
0.95073 | -84.2582
0.95459 | -84.5972
0.95845 | -85.2495
0.96232 | -85.8739
0.96618 | -85.7562
0.97005 | -85.3036
0.97391 | -85.4146
0.97778 | -86.0472
0.98164 | -85.9905
0.98551 | -85.9338
0.98937 | -85.8771
0.99324 | -85.8203
0.9971 | -85.7636
1.001 | -85.7069
1.0048 | -85.6502
1.0087 | -85.5935
1.0126 | -85.5368
1.0164 | -85.4801
1.0203 | -85.4233
1.0242 | -85.2582
1.028 | -84.9217
1.0319 | -84.893
1.0357 | -85.295
1.0396 | -86.037
1.0435 | -86.5533
1.0473 | -86.1134
1.0512 | -85.6599
1.0551 | -85.8891
1.0589 | -86.3117
1.0628 | -86.3579
1.0667 | -86.4323
1.0705 | -86.7894
1.0744 | -87.1783
1.0783 | -87.0037
1.0821 | -86.8934
1.086 | -86.8266
1.0899 | -86.3945
1.0937 | -86.2887
1.0976 | -86.7011
1.1014 | -87.1798
1.1053 | -87.4898
1.1092 | -87.8726
1.113 | -87.6768
1.1169 | -87.2517
1.1208 | -86.8052
1.1246 | -86.6299
1.1285 | -86.6805
1.1324 | -86.813
1.1362 | -86.9821
1.1401 | -87.1264
1.144 | -87.2189
1.1478 | -87.3392
1.1517 | -87.4662
1.1556 | -87.4901
1.1594 | -87.4233
1.1633 | -87.3165
1.1671 | -87.166
1.171 | -87.0695
1.1749 | -87.1563
1.1787 | -87.4015
1.1826 | -87.6298
1.1865 | -87.6452
1.1903 | -87.4038
1.1942 | -87.0663
1.1981 | -86.929
1.2019 | -87.1222
1.2058 | -87.4659
1.2097 | -87.7655
1.2135 | -88.0154
1.2174 | -88.1384
1.2213 | -88.0533
1.2251 | -87.9182
1.229 | -87.8808
1.2329 | -87.907
1.2367 | -87.8793
1.2406 | -87.6589
1.2444 | -87.3775
1.2483 | -87.396
1.2522 | -87.8149
1.256 | -88.2454
1.2599 | -88.2015
1.2638 | -87.8361
1.2676 | -87.605
1.2715 | -87.7054
1.2754 | -88.1387
1.2792 | -88.6387
1.2831 | -88.7335
1.287 | -88.355
1.2908 | -87.7595
1.2947 | -87.2614
1.2986 | -87.218
1.3024 | -87.6989
1.3063 | -88.2836
1.3101 | -88.3566
1.314 | -87.9673
1.3179 | -87.6788
1.3217 | -87.7034
1.3256 | -87.8855
1.3295 | -88.0254
1.3333 | -88.0647
1.3372 | -88.0968
1.3411 | -88.2226
1.3449 | -88.4136
1.3488 | -88.5231
1.3527 | -88.5568
1.3565 | -88.6892
1.3604 | -88.7835
1.3643 | -88.6347
1.3681 | -88.5227
1.372 | -88.7641
1.3758 | -89.1595
1.3797 | -89.2933
1.3836 | -89.1608
1.3874 | -88.969
1.3913 | -88.792
1.3952 | -88.7164
1.399 | -88.6835
1.4029 | -88.5046
1.4068 | -88.1945
1.4106 | -87.9953
1.4145 | -88.1219
1.4184 | -88.5739
1.4222 | -89.0544
1.4261 | -89.2711
1.43 | -89.3132
1.4338 | -89.4116
1.4377 | -89.5691
1.4415 | -89.6694
1.4454 | -89.6231
1.4493 | -89.3894
1.4531 | -89.1392
1.457 | -89.1909
1.4609 | -89.6019
1.4647 | -90.0168
1.4686 | -90.0697
1.4725 | -89.8538
1.4763 | -89.5486
1.4802 | -89.2982
1.4841 | -89.304
1.4879 | -89.4734
1.4918 | -89.517
1.4957 | -89.4673
1.4995 | -89.5163
1.5034 | -89.6574
1.5072 | -89.8394
1.5111 | -90.0017
1.515 | -90.0571
1.5188 | -90.081
1.5227 | -90.0623
1.5266 | -89.8054
1.5304 | -89.4626
1.5343 | -89.3317
1.5382 | -89.4594
1.542 | -89.7875
1.5459 | -90.2693
1.5498 | -90.7441
1.5536 | -90.9984
1.5575 | -90.9322
1.5614 | -90.5623
1.5652 | -90.1861
1.5691 | -90.0862
1.5729 | -90.1878
1.5768 | -90.2942
1.5807 | -90.4156
1.5845 | -90.5798
1.5884 | -90.656
1.5923 | -90.5091
1.5961 | -90.2031
1.6 | -89.9945
1.6039 | -89.9364
1.6077 | -89.723
1.6116 | -89.3433
1.6155 | -89.1568
1.6193 | -89.2605
1.6232 | -89.4284
1.6271 | -89.4703
1.6309 | -89.5098
1.6348 | -89.8434
1.6386 | -90.5302
1.6425 | -91.1962
1.6464 | -91.2613
1.6502 | -90.6897
1.6541 | -90.073
1.658 | -89.7509
1.6618 | -89.6046
1.6657 | -89.5416
1.6696 | -89.7867
1.6734 | -90.0092
1.6773 | -90.1855
1.6812 | -90.2245
1.685 | -90.2013
1.6889 | -90.2411
1.6928 | -90.3718
1.6966 | -90.4728
1.7005 | -90.4643
1.7043 | -90.4356
1.7082 | -90.5392
1.7121 | -90.8213
1.7159 | -91.2021
1.7198 | -91.5056
1.7237 | -91.6481
1.7275 | -91.6677
1.7314 | -91.6718
1.7353 | -91.723
1.7391 | -91.8215
1.743 | -91.8945
1.7469 | -91.8195
1.7507 | -91.5479
1.7546 | -91.1723
1.7585 | -90.8751
1.7623 | -90.8311
1.7662 | -91.0567
1.77 | -91.3845
1.7739 | -91.5601
1.7778 | -91.4401
1.7816 | -91.1516
1.7855 | -90.929
1.7894 | -90.889
1.7932 | -90.9695
1.7971 | -91.0412
1.801 | -91.0308
1.8048 | -90.9597
1.8087 | -90.8628
1.8126 | -90.7728
1.8164 | -90.7197
1.8203 | -90.7377
1.8242 | -90.8504
1.828 | -91.0403
1.8319 | -91.2603
1.8357 | -91.4896
1.8396 | -91.6965
1.8435 | -91.7945
1.8473 | -91.6924
1.8512 | -91.4914
1.8551 | -91.3431
1.8589 | -91.2907
1.8628 | -91.2983
1.8667 | -91.3613
1.8705 | -91.5127
1.8744 | -91.7295
1.8783 | -91.8787
1.8821 | -91.8698
1.886 | -91.7583
1.8899 | -91.7174
1.8937 | -91.8227
1.8976 | -92.0093
1.9014 | -92.1028
1.9053 | -92.0138
1.9092 | -91.8353
1.913 | -91.7254
1.9169 | -91.7274
1.9208 | -91.7546
1.9246 | -91.658
1.9285 | -91.4317
1.9324 | -91.229
1.9362 | -91.2074
1.9401 | -91.3738
1.944 | -91.6244
1.9478 | -91.813
1.9517 | -91.8598
1.9556 | -91.8204
1.9594 | -91.8093
1.9633 | -91.9106
1.9671 | -92.1273
1.971 | -92.3672
1.9749 | -92.536
1.9787 | -92.5721
1.9826 | -92.5069
1.9865 | -92.4407
1.9903 | -92.4348
1.9942 | -92.4494
1.9981 | -92.3748
2.0019 | -92.1932
2.0058 | -92.0297
2.0097 | -92.0094
2.0135 | -92.1442
2.0174 | -92.2895
2.0213 | -92.2827
2.0251 | -92.1598
2.029 | -92.0674
2.0328 | -92.0725
2.0367 | -92.1571
2.0406 | -92.2763
2.0444 | -92.3954
2.0483 | -92.4516
2.0522 | -92.4088
2.056 | -92.289
2.0599 | -92.1226
2.0638 | -91.9948
2.0676 | -92.0082
2.0715 | -92.1732
2.0754 | -92.3825
2.0792 | -92.4972
2.0831 | -92.4986
2.087 | -92.4748
2.0908 | -92.5228
2.0947 | -92.6456
2.0986 | -92.777
2.1024 | -92.86
2.1063 | -92.8401
2.1101 | -92.7361
2.114 | -92.6792
2.1179 | -92.757
2.1217 | -92.9161
2.1256 | -93.063
2.1295 | -93.2005
2.1333 | -93.3398
2.1372 | -93.4308
2.1411 | -93.4533
2.1449 | -93.384
2.1488 | -93.1931
2.1527 | -92.959
2.1565 | -92.7658
2.1604 | -92.6266
2.1643 | -92.5364
2.1681 | -92.5323
2.172 | -92.6244
2.1758 | -92.7625
2.1797 | -92.9182
2.1836 | -93.0896
2.1874 | -93.2408
2.1913 | -93.3557
2.1952 | -93.4162
2.199 | -93.3797
2.2029 | -93.3013
2.2068 | -93.2786
2.2106 | -93.3151
2.2145 | -93.3549
2.2184 | -93.3998
2.2222 | -93.4679
2.2261 | -93.5498
2.23 | -93.6454
2.2338 | -93.7078
2.2377 | -93.7206
2.2415 | -93.625
2.2454 | -93.418
2.2493 | -93.1643
2.2531 | -92.9508
2.257 | -92.8282
2.2609 | -92.8464
2.2647 | -93.0192
2.2686 | -93.3131
2.2725 | -93.6695
2.2763 | -94.015
2.2802 | -94.3059
2.2841 | -94.4937
2.2879 | -94.5422
2.2918 | -94.4599
2.2957 | -94.2786
2.2995 | -94.0647
2.3034 | -93.8433
2.3072 | -93.6704
2.3111 | -93.5841
2.315 | -93.5707
2.3188 | -93.595
2.3227 | -93.623
2.3266 | -93.6412
2.3304 | -93.6656
2.3343 | -93.6898
2.3382 | -93.7113
2.342 | -93.7156
2.3459 | -93.698
2.3498 | -93.6596
2.3536 | -93.6083
2.3575 | -93.5497
2.3614 | -93.483
2.3652 | -93.4064
2.3691 | -93.3271
2.3729 | -93.2618
2.3768 | -93.2557
2.3807 | -93.3195
2.3845 | -93.4555
2.3884 | -93.6259
2.3923 | -93.7977
2.3961 | -93.9357
2.4 | -94.0029
2.4039 | -93.9873
2.4077 | -93.9228
2.4116 | -93.8666
2.4155 | -93.8539
2.4193 | -93.8618
2.4232 | -93.8591
2.4271 | -93.844
2.4309 | -93.87
2.4348 | -93.943
2.4386 | -94.0288
2.4425 | -94.0714
2.4464 | -94.0807
2.4502 | -94.0953
2.4541 | -94.132
2.458 | -94.1661
2.4618 | -94.184
2.4657 | -94.1864
2.4696 | -94.1595
2.4734 | -94.1049
2.4773 | -94.0338
2.4812 | -93.9703
2.485 | -93.9348
2.4889 | -93.9157
2.4928 | -93.9153
2.4966 | -93.9129
2.5005 | -93.8842
2.5043 | -93.8309
2.5082 | -93.7879
2.5121 | -93.7947
2.5159 | -93.858
2.5198 | -93.9651
2.5237 | -94.0874
2.5275 | -94.2076
2.5314 | -94.3021
2.5353 | -94.3377
2.5391 | -94.3381
2.543 | -94.354
2.5469 | -94.421
2.5507 | -94.553
2.5546 | -94.7235
2.5585 | -94.859
2.5623 | -94.917
2.5662 | -94.9093
2.57 | -94.8607
2.5739 | -94.8086
2.5778 | -94.7679
2.5816 | -94.7253
2.5855 | -94.6596
2.5894 | -94.5505
2.5932 | -94.3685
2.5971 | -94.1201
2.601 | -93.8692
2.6048 | -93.6831
2.6087 | -93.6279
2.6126 | -93.7337
2.6164 | -93.9566
2.6203 | -94.2249
2.6242 | -94.4491
2.628 | -94.5665
2.6319 | -94.5795
2.6357 | -94.5479
2.6396 | -94.527
2.6435 | -94.5462
2.6473 | -94.6082
2.6512 | -94.7138
2.6551 | -94.864
2.6589 | -95.045
2.6628 | -95.1822
2.6667 | -95.2363
2.6705 | -95.2202
2.6744 | -95.1774
2.6783 | -95.1382
2.6821 | -95.0908
2.686 | -95.017
2.6899 | -94.9262
2.6937 | -94.8163
2.6976 | -94.7029
2.7014 | -94.5932
2.7053 | -94.5173
2.7092 | -94.5024
2.713 | -94.5619
2.7169 | -94.6624
2.7208 | -94.7729
2.7246 | -94.8833
2.7285 | -94.9815
2.7324 | -95.0554
2.7362 | -95.0864
2.7401 | -95.044
2.744 | -94.9196
2.7478 | -94.7288
2.7517 | -94.5396
2.7556 | -94.4379
2.7594 | -94.4916
2.7633 | -94.6821
2.7671 | -94.9479
2.771 | -95.1934
2.7749 | -95.3471
2.7787 | -95.3832
2.7826 | -95.3538
2.7865 | -95.4205
2.7903 | -95.4968
2.7942 | -95.6116
2.7981 | -95.7424
2.8019 | -95.839
2.8058 | -95.8711
2.8097 | -95.8379
2.8135 | -95.7455
2.8174 | -95.6269
2.8213 | -95.5178
2.8251 | -95.4587
2.829 | -95.4499
2.8329 | -95.4848
2.8367 | -95.5462
2.8406 | -95.6113
2.8444 | -95.6554
2.8483 | -95.6744
2.8522 | -95.68
2.856 | -95.6396
2.8599 | -95.5614
2.8638 | -95.4491
2.8676 | -95.3364
2.8715 | -95.2428
2.8754 | -95.186
2.8792 | -95.182
2.8831 | -95.2149
2.887 | -95.2594
2.8908 | -95.3027
2.8947 | -95.3445
2.8986 | -95.3835
2.9024 | -95.4265
2.9063 | -95.5053
2.9101 | -95.602
2.914 | -95.7043
2.9179 | -95.7831
2.9217 | -95.8272
2.9256 | -95.8456
2.9295 | -95.8525
2.9333 | -95.8701
2.9372 | -95.9095
2.9411 | -95.9525
2.9449 | -95.9883
2.9488 | -96.018
2.9527 | -96.0406
2.9565 | -96.0704
2.9604 | -96.1062
2.9643 | -96.1572
2.9681 | -96.2032
2.972 | -96.2319
2.9758 | -96.1868
2.9797 | -96.0761
2.9836 | -95.8778
2.9874 | -95.5685
2.9913 | -95.1947
2.9952 | -94.8429
2.999 | -94.6147
3.0029 | -94.5767
3.0068 | -94.7275
3.0106 | -95.0302
3.0145 | -95.3963
3.0184 | -95.7422
3.0222 | -96.0048
3.0261 | -96.1911
3.03 | -96.2924
3.0338 | -96.3508
3.0377 | -96.4126
3.0415 | -96.4802
3.0454 | -96.5561
3.0493 | -96.6405
3.0531 | -96.711
3.057 | -96.7599
3.0609 | -96.7557
3.0647 | -96.7093
3.0686 | -96.6196
3.0725 | -96.5109
3.0763 | -96.4063
3.0802 | -96.347
3.0841 | -96.3363
3.0879 | -96.3351
3.0918 | -96.3102
3.0957 | -96.2445
3.0995 | -96.1283
3.1034 | -95.9947
3.1072 | -95.8883
3.1111 | -95.8481
3.115 | -95.8796
3.1188 | -95.9807
3.1227 | -96.1192
3.1266 | -96.2524
3.1304 | -96.3429
3.1343 | -96.3612
3.1382 | -96.3067
3.142 | -96.19
3.1459 | -96.0365
3.1498 | -95.9154
3.1536 | -95.8651
3.1575 | -95.876
3.1614 | -95.9432
3.1652 | -96.0375
3.1691 | -96.1435
3.1729 | -96.2432
3.1768 | -96.3316
3.1807 | -96.4017
3.1845 | -96.4636
3.1884 | -96.5012
3.1923 | -96.5105
3.1961 | -96.4925
3.2 | -96.4404
3.2039 | -96.3666
3.2077 | -96.2958
3.2116 | -96.2296
3.2155 | -96.1828
3.2193 | -96.1719
3.2232 | -96.2191
3.2271 | -96.3378
3.2309 | -96.533
3.2348 | -96.779
3.2386 | -97.0289
3.2425 | -97.2092
3.2464 | -97.2762
3.2502 | -97.2155
3.2541 | -97.0638
3.258 | -96.8766
3.2618 | -96.707
3.2657 | -96.5801
3.2696 | -96.5112
3.2734 | -96.4817
3.2773 | -96.4705
3.2812 | -96.4433
3.285 | -96.3998
3.2889 | -96.3505
3.2928 | -96.3322
3.2966 | -96.3592
3.3005 | -96.4233
3.3043 | -96.4982
3.3082 | -96.5597
3.3121 | -96.5952
3.3159 | -96.6159
3.3198 | -96.6251
3.3237 | -96.6389
3.3275 | -96.6583
3.3314 | -96.681
3.3353 | -96.6979
3.3391 | -96.7167
3.343 | -96.7214
3.3469 | -96.7174
3.3507 | -96.6939
3.3546 | -96.6477
3.3585 | -96.5813
3.3623 | -96.5129
3.3662 | -96.4558
3.37 | -96.4331
3.3739 | -96.449
3.3778 | -96.5022
3.3816 | -96.584
3.3855 | -96.6874
3.3894 | -96.7872
3.3932 | -96.8664
3.3971 | -96.9087
3.401 | -96.912
3.4048 | -96.8921
3.4087 | -96.8626
3.4126 | -96.8378
3.4164 | -96.8363
3.4203 | -96.8624
3.4242 | -96.9237
3.428 | -96.9974
3.4319 | -97.0743
3.4357 | -97.1419
3.4396 | -97.1917
3.4435 | -97.2237
3.4473 | -97.2511
3.4512 | -97.2631
3.4551 | -97.2555
3.4589 | -97.2212
3.4628 | -97.1534
3.4667 | -97.0734
3.4705 | -96.9997
3.4744 | -96.9338
3.4783 | -96.887
3.4821 | -96.8699
3.486 | -96.8783
3.4899 | -96.9189
3.4937 | -96.9808
3.4976 | -97.06
3.5014 | -97.1513
3.5053 | -97.2492
3.5092 | -97.3506
3.513 | -97.4494
3.5169 | -97.5423
3.5208 | -97.6204
3.5246 | -97.6772
3.5285 | -97.7058
3.5324 | -97.7097
3.5362 | -97.6791
3.5401 | -97.6203
3.544 | -97.5499
3.5478 | -97.4772
3.5517 | -97.414
3.5556 | -97.3642
3.5594 | -97.3166
3.5633 | -97.266
3.5671 | -97.2104
3.571 | -97.1491
3.5749 | -97.0856
3.5787 | -97.0345
3.5826 | -97.0097
3.5865 | -97.0241
3.5903 | -97.0729
3.5942 | -97.1439
3.5981 | -97.2113
3.6019 | -97.2637
3.6058 | -97.2898
3.6097 | -97.3007
3.6135 | -97.3153
3.6174 | -97.3457
3.6213 | -97.3785
3.6251 | -97.3986
3.629 | -97.4035
3.6328 | -97.3907
3.6367 | -97.374
3.6406 | -97.3635
3.6444 | -97.3691
3.6483 | -97.3927
3.6522 | -97.4186
3.656 | -97.445
3.6599 | -97.471
3.6638 | -97.4999
3.6676 | -97.5499
3.6715 | -97.612
3.6754 | -97.6748
3.6792 | -97.7299
3.6831 | -97.764
3.687 | -97.753
3.6908 | -97.7125
3.6947 | -97.6649
3.6986 | -97.6271
3.7024 | -97.6088
3.7063 | -97.6201
3.7101 | -97.6655
3.714 | -97.7255
3.7179 | -97.7814
3.7217 | -97.83
3.7256 | -97.8344
3.7295 | -97.7899
3.7333 | -97.7244
3.7372 | -97.6498
3.7411 | -97.5937
3.7449 | -97.567
3.7488 | -97.5574
3.7527 | -97.5529
3.7565 | -97.5301
3.7604 | -97.4934
3.7643 | -97.4523
3.7681 | -97.439
3.772 | -97.478
3.7758 | -97.5764
3.7797 | -97.7179
3.7836 | -97.8965
3.7874 | -98.0706
3.7913 | -98.1884
3.7952 | -98.2435
3.799 | -98.2302
3.8029 | -98.1797
3.8068 | -98.1189
3.8106 | -98.0739
3.8145 | -98.0878
3.8184 | -98.154
3.8222 | -98.2779
3.8261 | -98.4365
3.83 | -98.6147
3.8338 | -98.7703
3.8377 | -98.8819
3.8415 | -98.9315
3.8454 | -98.9217
3.8493 | -98.8406
3.8531 | -98.704
3.857 | -98.5365
3.8609 | -98.354
3.8647 | -98.1739
3.8686 | -98.0056
3.8725 | -97.8526
3.8763 | -97.7332
3.8802 | -97.6572
3.8841 | -97.6319
3.8879 | -97.6492
3.8918 | -97.7017
3.8957 | -97.769
3.8995 | -97.8506
3.9034 | -97.9183
3.9072 | -97.9942
3.9111 | -98.079
3.915 | -98.1558
3.9188 | -98.2189
3.9227 | -98.2685
3.9266 | -98.3053
3.9304 | -98.3433
3.9343 | -98.3878
3.9382 | -98.4323
3.942 | -98.4783
3.9459 | -98.5241
3.9498 | -98.5589
3.9536 | -98.5903
3.9575 | -98.6264
3.9614 | -98.663
3.9652 | -98.7033
3.9691 | -98.7383
3.9729 | -98.7429
3.9768 | -98.723
3.9807 | -98.6721
3.9845 | -98.5961
3.9884 | -98.5016
3.9923 | -98.4047
3.9961 | -98.3182
4 | -98.2607
4.0039 | -98.2473
4.0077 | -98.269
4.0116 | -98.3161
4.0155 | -98.3796
4.0193 | -98.4323
4.0232 | -98.4588
4.0271 | -98.4603
4.0309 | -98.4376
4.0348 | -98.3987
4.0386 | -98.3574
4.0425 | -98.3158
4.0464 | -98.28
4.0502 | -98.2571
4.0541 | -98.2552
4.058 | -98.2723
4.0618 | -98.314
4.0657 | -98.3735
4.0696 | -98.4417
4.0734 | -98.5172
4.0773 | -98.584
4.0812 | -98.6301
4.085 | -98.6562
4.0889 | -98.6719
4.0928 | -98.683
4.0966 | -98.6937
4.1005 | -98.7084
4.1043 | -98.7325
4.1082 | -98.7525
4.1121 | -98.7761
4.1159 | -98.784
4.1198 | -98.7761
4.1237 | -98.7566
4.1275 | -98.7163
4.1314 | -98.6778
4.1353 | -98.6462
4.1391 | -98.6156
4.143 | -98.6104
4.1469 | -98.62
4.1507 | -98.6418
4.1546 | -98.6589
4.1585 | -98.6676
4.1623 | -98.6582
4.1662 | -98.6302
4.17 | -98.5926
4.1739 | -98.5642
4.1778 | -98.5504
4.1816 | -98.5631
4.1855 | -98.5908
4.1894 | -98.6235
4.1932 | -98.6467
4.1971 | -98.6452
4.201 | -98.6164
4.2048 | -98.5835
4.2087 | -98.5382
4.2126 | -98.4992
4.2164 | -98.4693
4.2203 | -98.4619
4.2242 | -98.4804
4.228 | -98.521
4.2319 | -98.5795
4.2357 | -98.6381
4.2396 | -98.6761
4.2435 | -98.6863
4.2473 | -98.678
4.2512 | -98.6562
4.2551 | -98.6385
4.2589 | -98.6551
4.2628 | -98.7038
4.2667 | -98.7805
4.2705 | -98.858
4.2744 | -98.9364
4.2783 | -99.0065
4.2821 | -99.0572
4.286 | -99.0886
4.2899 | -99.111
4.2937 | -99.1172
4.2976 | -99.1079
4.3014 | -99.0883
4.3053 | -99.0682
4.3092 | -99.044
4.313 | -99.0293
4.3169 | -99.0323
4.3208 | -99.0495
4.3246 | -99.0772
4.3285 | -99.1121
4.3324 | -99.147
4.3362 | -99.1828
4.3401 | -99.2054
4.344 | -99.2072
4.3478 | -99.1937
4.3517 | -99.1651
4.3556 | -99.1139
4.3594 | -99.0586
4.3633 | -99.0045
4.3671 | -98.9461
4.371 | -98.8932
4.3749 | -98.8457
4.3787 | -98.8055
4.3826 | -98.7898
4.3865 | -98.8034
4.3903 | -98.8478
4.3942 | -98.9256
4.3981 | -99.0151
4.4019 | -99.0926
4.4058 | -99.1588
4.4097 | -99.1993
4.4135 | -99.211
4.4174 | -99.1937
4.4213 | -99.1604
4.4251 | -99.127
4.429 | -99.0974
4.4329 | -99.0978
4.4367 | -99.1167
4.4406 | -99.1589
4.4444 | -99.2075
4.4483 | -99.2423
4.4522 | -99.2632
4.456 | -99.2637
4.4599 | -99.2767
4.4638 | -99.2969
4.4676 | -99.3348
4.4715 | -99.3858
4.4754 | -99.4428
4.4792 | -99.5118
4.4831 | -99.5755
4.487 | -99.6312
4.4908 | -99.6605
4.4947 | -99.6585
4.4986 | -99.6405
4.5024 | -99.6098
4.5063 | -99.5859
4.5101 | -99.574
4.514 | -99.5813
4.5179 | -99.6185
4.5217 | -99.6797
4.5256 | -99.7658
4.5295 | -99.8665
4.5333 | -99.9776
4.5372 | -100.102
4.5411 | -100.2336
4.5449 | -100.3582
4.5488 | -100.454
4.5527 | -100.5109
4.5565 | -100.541
4.5604 | -100.5429
4.5643 | -100.5307
4.5681 | -100.502
4.572 | -100.4754
4.5758 | -100.4569
4.5797 | -100.4426
4.5836 | -100.4347
4.5874 | -100.4404
4.5913 | -100.4298
4.5952 | -100.3967
4.599 | -100.3409
4.6029 | -100.2589
4.6068 | -100.1461
4.6106 | -100.012
4.6145 | -99.8666
4.6184 | -99.7325
4.6222 | -99.6208
4.6261 | -99.5352
4.63 | -99.4738
4.6338 | -99.4459
4.6377 | -99.4489
4.6415 | -99.4778
4.6454 | -99.5239
4.6493 | -99.583
4.6531 | -99.6485
4.657 | -99.7103
4.6609 | -99.7541
4.6647 | -99.7879
4.6686 | -99.8023
4.6725 | -99.7961
4.6763 | -99.7807
4.6802 | -99.7553
4.6841 | -99.7197
4.6879 | -99.6839
4.6918 | -99.6497
4.6957 | -99.6181
4.6995 | -99.6008
4.7034 | -99.5983
4.7072 | -99.6049
4.7111 | -99.6287
4.715 | -99.6591
4.7188 | -99.7018
4.7227 | -99.7545
4.7266 | -99.8072
4.7304 | -99.851
4.7343 | -99.88
4.7382 | -99.8931
4.742 | -99.8815
4.7459 | -99.8548
4.7498 | -99.8173
4.7536 | -99.7879
4.7575 | -99.7765
4.7614 | -99.7777
4.7652 | -99.7896
4.7691 | -99.8139
4.7729 | -99.8405
4.7768 | -99.8729
4.7807 | -99.9211
4.7845 | -99.9823
4.7884 | -100.0591
4.7923 | -100.1375
4.7961 | -100.2239
4.8 | -100.3068
4.8039 | -100.3799
4.8077 | -100.4358
4.8116 | -100.4745
4.8155 | -100.4948
4.8193 | -100.4884
4.8232 | -100.4718
4.8271 | -100.4444
4.8309 | -100.3977
4.8348 | -100.3525
4.8386 | -100.3157
4.8425 | -100.2805
4.8464 | -100.253
4.8502 | -100.2316
4.8541 | -100.2228
4.858 | -100.2184
4.8618 | -100.2218
4.8657 | -100.2288
4.8696 | -100.2316
4.8734 | -100.2406
4.8773 | -100.2397
4.8812 | -100.2339
4.885 | -100.2104
4.8889 | -100.1789
4.8928 | -100.1359
4.8966 | -100.0806
4.9005 | -100.0349
4.9043 | -100.0027
4.9082 | -99.9801
4.9121 | -99.9669
4.9159 | -99.9694
4.9198 | -99.9894
4.9237 | -100.0312
4.9275 | -100.0863
4.9314 | -100.1686
4.9353 | -100.2708
4.9391 | -100.3982
4.943 | -100.5308
4.9469 | -100.6594
4.9507 | -100.7759
4.9546 | -100.8569
4.9585 | -100.9033
4.9623 | -100.8703
4.9662 | -100.7522
4.97 | -100.5954
4.9739 | -100.383
4.9778 | -100.1311
4.9816 | -99.8542
4.9855 | -99.5888
4.9894 | -99.3574
4.9932 | -99.1859
4.9971 | -99.0967
5.001 | -99.0983
5.0048 | -99.1863
5.0087 | -99.3447
5.0126 | -99.5528
5.0164 | -99.7536
5.0203 | -99.9339
5.0242 | -100.095
5.028 | -100.2271
5.0319 | -100.3166
5.0357 | -100.3571
5.0396 | -100.3882
5.0435 | -100.4073
5.0473 | -100.3918
5.0512 | -100.3855
5.0551 | -100.401
5.0589 | -100.4468
5.0628 | -100.5102
5.0667 | -100.5891
5.0705 | -100.6782
5.0744 | -100.7727
5.0783 | -100.8626
5.0821 | -100.9419
5.086 | -101.0073
5.0899 | -101.0491
5.0937 | -101.0701
5.0976 | -101.0709
5.1014 | -101.0502
5.1053 | -101.0095
5.1092 | -100.9637
5.113 | -100.9233
5.1169 | -100.8806
5.1208 | -100.8303
5.1246 | -100.7813
5.1285 | -100.7408
5.1324 | -100.7138
5.1362 | -100.6964
5.1401 | -100.6841
5.144 | -100.6832
5.1478 | -100.688
5.1517 | -100.6943
5.1556 | -100.6964
5.1594 | -100.7001
5.1633 | -100.7013
5.1672 | -100.699
5.171 | -100.6887
5.1749 | -100.6791
5.1787 | -100.6674
5.1826 | -100.6502
5.1865 | -100.6304
5.1903 | -100.6008
5.1942 | -100.5684
5.1981 | -100.5405
5.2019 | -100.5162
5.2058 | -100.4989
5.2097 | -100.491
5.2135 | -100.4869
5.2174 | -100.4919
5.2213 | -100.4999
5.2251 | -100.5015
5.229 | -100.5055
5.2329 | -100.5079
5.2367 | -100.5105
5.2406 | -100.504
5.2444 | -100.4947
5.2483 | -100.4994
5.2522 | -100.5128
5.256 | -100.5352
5.2599 | -100.576
5.2638 | -100.6356
5.2676 | -100.7049
5.2715 | -100.7778
5.2754 | -100.8404
5.2792 | -100.8867
5.2831 | -100.915
5.287 | -100.9133
5.2908 | -100.8857
5.2947 | -100.8401
5.2986 | -100.7849
5.3024 | -100.7286
5.3063 | -100.6744
5.3101 | -100.628
5.314 | -100.6012
5.3179 | -100.5988
5.3217 | -100.6232
5.3256 | -100.6723
5.3295 | -100.7294
5.3333 | -100.8001
5.3372 | -100.8704
5.3411 | -100.9228
5.3449 | -100.9564
5.3488 | -100.9775
5.3527 | -100.9795
5.3565 | -100.9639
5.3604 | -100.9294
5.3643 | -100.8835
5.3681 | -100.8395
5.372 | -100.7962
5.3758 | -100.7615
5.3797 | -100.7354
5.3836 | -100.7272
5.3874 | -100.7372
5.3913 | -100.7554
5.3952 | -100.7766
5.399 | -100.8073
5.4029 | -100.8386
5.4068 | -100.8494
5.4106 | -100.8453
5.4145 | -100.8325
5.4184 | -100.8118
5.4222 | -100.7912
5.4261 | -100.7654
5.43 | -100.7365
5.4338 | -100.7094
5.4377 | -100.6879
5.4415 | -100.6745
5.4454 | -100.66
5.4493 | -100.6458
5.4531 | -100.6381
5.457 | -100.6267
5.4609 | -100.6177
5.4647 | -100.6149
5.4686 | -100.6099
5.4725 | -100.6093
5.4763 | -100.6106
5.4802 | -100.6055
5.4841 | -100.6
5.4879 | -100.5971
5.4918 | -100.5905
5.4957 | -100.5726
5.4995 | -100.5573
5.5034 | -100.5515
5.5072 | -100.5621
5.5111 | -100.5902
5.515 | -100.6395
5.5188 | -100.702
5.5227 | -100.7688
5.5266 | -100.8324
5.5304 | -100.8919
5.5343 | -100.9332
5.5382 | -100.9572
5.542 | -100.9691
5.5459 | -100.9662
5.5498 | -100.9562
5.5536 | -100.9316
5.5575 | -100.9029
5.5614 | -100.8733
5.5652 | -100.8556
5.5691 | -100.8553
5.5729 | -100.9025
5.5768 | -100.9299
5.5807 | -100.9732
5.5845 | -101.0223
5.5884 | -101.0737
5.5923 | -101.1314
5.5961 | -101.1795
5.6 | -101.2117
5.6039 | -101.2337
5.6077 | -101.2504
5.6116 | -101.2646
5.6155 | -101.2665
5.6193 | -101.2606
5.6232 | -101.2657
5.6271 | -101.2676
5.6309 | -101.2778
5.6348 | -101.3068
5.6386 | -101.3456
5.6425 | -101.3885
5.6464 | -101.4372
5.6502 | -101.4776
5.6541 | -101.5229
5.658 | -101.5711
5.6618 | -101.6131
5.6657 | -101.6468
5.6696 | -101.6671
5.6734 | -101.6826
5.6773 | -101.6863
5.6812 | -101.679
5.685 | -101.67
5.6889 | -101.647
5.6928 | -101.609
5.6966 | -101.5685
5.7005 | -101.5362
5.7043 | -101.5147
5.7082 | -101.5005
5.7121 | -101.4855
5.7159 | -101.4841
5.7198 | -101.4924
5.7237 | -101.4992
5.7275 | -101.516
5.7314 | -101.539
5.7353 | -101.5575
5.7391 | -101.5731
5.743 | -101.5795
5.7469 | -101.5837
5.7507 | -101.5802
5.7546 | -101.5692
5.7585 | -101.5566
5.7623 | -101.5371
5.7662 | -101.515
5.77 | -101.488
5.7739 | -101.4653
5.7778 | -101.4508
5.7816 | -101.4399
5.7855 | -101.4252
5.7894 | -101.4081
5.7932 | -101.3963
5.7971 | -101.3989
5.801 | -101.4161
5.8048 | -101.436
5.8087 | -101.4613
5.8126 | -101.5036
5.8164 | -101.5418
5.8203 | -101.5785
5.8242 | -101.6211
5.828 | -101.6689
5.8319 | -101.7158
5.8357 | -101.7604
5.8396 | -101.8123
5.8435 | -101.8565
5.8473 | -101.8872
5.8512 | -101.9076
5.8551 | -101.923
5.8589 | -101.9322
5.8628 | -101.9229
5.8667 | -101.9044
5.8705 | -101.8816
5.8744 | -101.8573
5.8783 | -101.8332
5.8821 | -101.8163
5.886 | -101.8031
5.8899 | -101.7897
5.8937 | -101.7887
5.8976 | -101.782
5.9014 | -101.7705
5.9053 | -101.7648
5.9092 | -101.7506
5.913 | -101.7285
5.9169 | -101.6966
5.9208 | -101.6635
5.9246 | -101.6393
5.9285 | -101.6136
5.9324 | -101.5973
5.9362 | -101.5919
5.9401 | -101.5861
5.944 | -101.5828
5.9478 | -101.5942
5.9517 | -101.6276
5.9556 | -101.6677
5.9594 | -101.7114
5.9633 | -101.7591
5.9672 | -101.8057
5.971 | -101.8447
5.9749 | -101.8817
5.9787 | -101.9158
5.9826 | -101.9359
5.9865 | -101.9543
5.9903 | -101.9645
5.9942 | -101.968
5.9981 | -101.9699
6.0019 | -101.9658
6.0058 | -101.9574
6.0097 | -101.9356
6.0135 | -101.9058
6.0174 | -101.879
6.0213 | -101.8525
6.0251 | -101.8246
6.029 | -101.801
6.0328 | -101.7709
6.0367 | -101.7308
6.0406 | -101.6867
6.0444 | -101.6441
6.0483 | -101.6066
6.0522 | -101.572
6.056 | -101.5515
6.0599 | -101.5505
6.0638 | -101.5656
6.0676 | -101.6024
6.0715 | -101.6637
6.0754 | -101.7406
6.0792 | -101.8224
6.0831 | -101.9079
6.087 | -101.989
6.0908 | -102.0605
6.0947 | -102.112
6.0986 | -102.1454
6.1024 | -102.1515
6.1063 | -102.1416
6.1101 | -102.1096
6.114 | -102.0636
6.1179 | -102.0179
6.1217 | -101.9725
6.1256 | -101.92
6.1295 | -101.8648
6.1333 | -101.8098
6.1372 | -101.7619
6.1411 | -101.7179
6.1449 | -101.6814
6.1488 | -101.649
6.1527 | -101.623
6.1565 | -101.6118
6.1604 | -101.6131
6.1643 | -101.6215
6.1681 | -101.6347
6.172 | -101.6546
6.1758 | -101.668
6.1797 | -101.682
6.1836 | -101.6917
6.1874 | -101.6935
6.1913 | -101.6875
6.1952 | -101.6737
6.199 | -101.6408
6.2029 | -101.6018
6.2068 | -101.5637
6.2106 | -101.5181
6.2145 | -101.4785
6.2184 | -101.4526
6.2222 | -101.4314
6.2261 | -101.4181
6.23 | -101.4187
6.2338 | -101.4254
6.2377 | -101.4407
6.2415 | -101.4633
6.2454 | -101.4891
6.2493 | -101.5139
6.2531 | -101.5357
6.257 | -101.5538
6.2609 | -101.5739
6.2647 | -101.5979
6.2686 | -101.6178
6.2725 | -101.64
6.2763 | -101.6647
6.2802 | -101.6933
6.2841 | -101.7264
6.2879 | -101.7605
6.2918 | -101.7958
6.2957 | -101.8343
6.2995 | -101.8587
6.3034 | -101.871
6.3072 | -101.87
6.3111 | -101.8543
6.315 | -101.8273
6.3188 | -101.8012
6.3227 | -101.7647
6.3266 | -101.7155
6.3304 | -101.669
6.3343 | -101.6297
6.3382 | -101.603
6.342 | -101.5857
6.3459 | -101.5786
6.3498 | -101.5817
6.3536 | -101.5969
6.3575 | -101.6217
6.3614 | -101.6539
6.3652 | -101.6898
6.3691 | -101.7281
6.3729 | -101.7548
6.3768 | -101.7707
6.3807 | -101.7789
6.3845 | -101.7857
6.3884 | -101.7889
6.3923 | -101.7886
6.3961 | -101.7969
6.4 | -101.8028
6.4039 | -101.8095
6.4077 | -101.817
6.4116 | -101.8245
6.4155 | -101.8332
6.4193 | -101.8477
6.4232 | -101.867
6.4271 | -101.8841
6.4309 | -101.9054
6.4348 | -101.935
6.4386 | -101.9773
6.4425 | -102.0293
6.4464 | -102.0879
6.4502 | -102.1457
6.4541 | -102.2051
6.458 | -102.2681
6.4618 | -102.329
6.4657 | -102.3777
6.4696 | -102.4138
6.4734 | -102.4406
6.4773 | -102.4545
6.4812 | -102.4501
6.485 | -102.4391
6.4889 | -102.4239
6.4928 | -102.4017
6.4966 | -102.3809
6.5005 | -102.3646
6.5043 | -102.3489
6.5082 | -102.3359
6.5121 | -102.3217
6.5159 | -102.3157
6.5198 | -102.3103
6.5237 | -102.3001
6.5275 | -102.281
6.5314 | -102.2614
6.5353 | -102.2448
6.5391 | -102.2347
6.543 | -102.233
6.5469 | -102.2396
6.5507 | -102.2442
6.5546 | -102.2527
6.5585 | -102.2747
6.5623 | -102.3008
6.5662 | -102.3368
6.57 | -102.3737
6.5739 | -102.4225
6.5778 | -102.4725
6.5816 | -102.5253
6.5855 | -102.5789
6.5894 | -102.6237
6.5932 | -102.6568
6.5971 | -102.6869
6.601 | -102.7003
6.6048 | -102.7035
6.6087 | -102.7062
6.6126 | -102.7023
6.6164 | -102.6975
6.6203 | -102.6952
6.6242 | -102.7009
6.628 | -102.7058
6.6319 | -102.7084
6.6357 | -102.7093
6.6396 | -102.7106
6.6435 | -102.7078
6.6473 | -102.7056
6.6512 | -102.701
6.6551 | -102.6896
6.6589 | -102.676
6.6628 | -102.6517
6.6667 | -102.6254
6.6705 | -102.5994
6.6744 | -102.5769
6.6783 | -102.5508
6.6821 | -102.5302
6.686 | -102.5143
6.6899 | -102.5089
6.6937 | -102.5015
6.6976 | -102.4919
6.7014 | -102.495
6.7053 | -102.5067
6.7092 | -102.5238
6.713 | -102.5489
6.7169 | -102.5799
6.7208 | -102.61
6.7246 | -102.6315
6.7285 | -102.6497
6.7324 | -102.6628
6.7362 | -102.6644
6.7401 | -102.6525
6.744 | -102.6333
6.7478 | -102.6045
6.7517 | -102.5669
6.7556 | -102.5291
6.7594 | -102.4951
6.7633 | -102.4636
6.7671 | -102.4325
6.771 | -102.4026
6.7749 | -102.3862
6.7787 | -102.3838
6.7826 | -102.3845
6.7865 | -102.3881
6.7903 | -102.4023
6.7942 | -102.4184
6.7981 | -102.4369
6.8019 | -102.4536
6.8058 | -102.4707
6.8097 | -102.4817
6.8135 | -102.4892
6.8174 | -102.4957
6.8213 | -102.5032
6.8251 | -102.5131
6.829 | -102.5217
6.8328 | -102.53
6.8367 | -102.5496
6.8406 | -102.5737
6.8444 | -102.5983
6.8483 | -102.6197
6.8522 | -102.6392
6.856 | -102.6573
6.8599 | -102.6701
6.8638 | -102.6757
6.8676 | -102.6799
6.8715 | -102.6825
6.8754 | -102.6848
6.8792 | -102.6901
6.8831 | -102.6959
6.887 | -102.7092
6.8908 | -102.7308
6.8947 | -102.7554
6.8986 | -102.7734
6.9024 | -102.7936
6.9063 | -102.816
6.9101 | -102.8283
6.914 | -102.84
6.9179 | -102.8493
6.9217 | -102.8451
6.9256 | -102.8286
6.9295 | -102.8023
6.9333 | -102.7722
6.9372 | -102.7381
6.9411 | -102.6933
6.9449 | -102.6474
6.9488 | -102.5912
6.9527 | -102.5418
6.9565 | -102.4905
6.9604 | -102.4338
6.9643 | -102.3738
6.9681 | -102.316
6.972 | -102.2579
6.9758 | -102.1983
6.9797 | -102.1453
6.9836 | -102.1063
6.9874 | -102.0809
6.9913 | -102.0622
6.9952 | -102.0544
6.999 | -102.065
7.0029 | -102.095
7.0068 | -102.133
7.0106 | -102.179
7.0145 | -102.2377
7.0184 | -102.2967
7.0222 | -102.3602
7.0261 | -102.4295
7.03 | -102.4967
7.0338 | -102.5587
7.0377 | -102.6128
7.0415 | -102.6584
7.0454 | -102.699
7.0493 | -102.73
7.0531 | -102.7593
7.057 | -102.7698
7.0609 | -102.7738
7.0647 | -102.7761
7.0686 | -102.7765
7.0725 | -102.7714
7.0763 | -102.7648
7.0802 | -102.757
7.0841 | -102.7438
7.0879 | -102.7219
7.0918 | -102.7017
7.0957 | -102.6944
7.0995 | -102.6881
7.1034 | -102.6774
7.1072 | -102.6735
7.1111 | -102.6737
7.115 | -102.6732
7.1188 | -102.6774
7.1227 | -102.6924
7.1266 | -102.7095
7.1304 | -102.7321
7.1343 | -102.7619
7.1382 | -102.7952
7.142 | -102.8328
7.1459 | -102.8717
7.1498 | -102.9154
7.1536 | -102.9627
7.1575 | -103.0096
7.1614 | -103.053
7.1652 | -103.093
7.1691 | -103.1297
7.1729 | -103.1656
7.1768 | -103.1969
7.1807 | -103.2159
7.1845 | -103.2223
7.1884 | -103.2138
7.1923 | -103.1927
7.1961 | -103.1624
7.2 | -103.1157
7.2039 | -103.0633
7.2077 | -103.0086
7.2116 | -102.9535
7.2155 | -102.9086
7.2193 | -102.8809
7.2232 | -102.856
7.2271 | -102.8399
7.2309 | -102.8357
7.2348 | -102.8398
7.2386 | -102.8483
7.2425 | -102.8668
7.2464 | -102.8884
7.2502 | -102.9006
7.2541 | -102.9169
7.258 | -102.9295
7.2618 | -102.9373
7.2657 | -102.943
7.2696 | -102.9466
7.2734 | -102.9503
7.2773 | -102.9453
7.2812 | -102.9318
7.285 | -102.9222
7.2889 | -102.9178
7.2928 | -102.9124
7.2966 | -102.9089
7.3005 | -102.9087
7.3043 | -102.9096
7.3082 | -102.9156
7.3121 | -102.9277
7.3159 | -102.9378
7.3198 | -102.9488
7.3237 | -102.9599
7.3275 | -102.9671
7.3314 | -102.977
7.3353 | -102.9883
7.3391 | -103.0009
7.343 | -103.0145
7.3469 | -103.0274
7.3507 | -103.0327
7.3546 | -103.0411
7.3585 | -103.0545
7.3623 | -103.0678
7.3662 | -103.0809
7.37 | -103.0971
7.3739 | -103.1182
7.3778 | -103.1413
7.3816 | -103.1636
7.3855 | -103.1816
7.3894 | -103.205
7.3932 | -103.2275
7.3971 | -103.2361
7.401 | -103.2346
7.4048 | -103.2299
7.4087 | -103.217
7.4126 | -103.1944
7.4164 | -103.1664
7.4203 | -103.1294
7.4242 | -103.0823
7.428 | -103.0433
7.4319 | -103.0077
7.4357 | -102.9674
7.4396 | -102.93
7.4435 | -102.904
7.4473 | -102.8824
7.4512 | -102.8628
7.4551 | -102.8558
7.4589 | -102.8603
7.4628 | -102.8741
7.4667 | -102.8981
7.4705 | -102.9338
7.4744 | -102.9725
7.4783 | -103.0226
7.4821 | -103.0811
7.486 | -103.1437
7.4899 | -103.2077
7.4937 | -103.2687
7.4976 | -103.3241
7.5014 | -103.3724
7.5053 | -103.4077
7.5092 | -103.4371
7.513 | -103.4561
7.5169 | -103.4613
7.5208 | -103.4565
7.5246 | -103.4456
7.5285 | -103.4267
7.5324 | -103.4116
7.5362 | -103.4032
7.5401 | -103.3835
7.544 | -103.3629
7.5478 | -103.3518
7.5517 | -103.3399
7.5556 | -103.3239
7.5594 | -103.31
7.5633 | -103.2932
7.5671 | -103.2734
7.571 | -103.2557
7.5749 | -103.242
7.5787 | -103.2227
7.5826 | -103.203
7.5865 | -103.1931
7.5903 | -103.1821
7.5942 | -103.1736
7.5981 | -103.1731
7.6019 | -103.1796
7.6058 | -103.1993
7.6097 | -103.2266
7.6135 | -103.2592
7.6174 | -103.3034
7.6213 | -103.3508
7.6251 | -103.398
7.629 | -103.4479
7.6329 | -103.4973
7.6367 | -103.5435
7.6406 | -103.5835
7.6444 | -103.6071
7.6483 | -103.6153
7.6522 | -103.6192
7.656 | -103.6146
7.6599 | -103.5948
7.6638 | -103.5699
7.6676 | -103.5421
7.6715 | -103.5083
7.6754 | -103.4789
7.6792 | -103.4449
7.6831 | -103.4079
7.687 | -103.3695
7.6908 | -103.3375
7.6947 | -103.3046
7.6986 | -103.2699
7.7024 | -103.2402
7.7063 | -103.2185
7.7101 | -103.1977
7.714 | -103.1875
7.7179 | -103.1854
7.7217 | -103.1873
7.7256 | -103.197
7.7295 | -103.2101
7.7333 | -103.2272
7.7372 | -103.25
7.7411 | -103.2659
7.7449 | -103.276
7.7488 | -103.2866
7.7527 | -103.291
7.7565 | -103.2959
7.7604 | -103.2984
7.7643 | -103.2944
7.7681 | -103.2863
7.772 | -103.2816
7.7758 | -103.2714
7.7797 | -103.2604
7.7836 | -103.2525
7.7874 | -103.2431
7.7913 | -103.2372
7.7952 | -103.239
7.799 | -103.2462
7.8029 | -103.2642
7.8068 | -103.2757
7.8106 | -103.293
7.8145 | -103.3137
7.8184 | -103.338
7.8222 | -103.3598
7.8261 | -103.3853
7.83 | -103.4029
7.8338 | -103.42
7.8377 | -103.4428
7.8415 | -103.4601
7.8454 | -103.47
7.8493 | -103.4739
7.8531 | -103.4778
7.857 | -103.483
7.8609 | -103.4843
7.8647 | -103.479
7.8686 | -103.4761
7.8725 | -103.4729
7.8763 | -103.4754
7.8802 | -103.4825
7.8841 | -103.4911
7.8879 | -103.5046
7.8918 | -103.5278
7.8957 | -103.5527
7.8995 | -103.5772
7.9034 | -103.606
7.9072 | -103.6334
7.9111 | -103.6586
7.915 | -103.6858
7.9188 | -103.7147
7.9227 | -103.7405
7.9266 | -103.7627
7.9304 | -103.7832
7.9343 | -103.8028
7.9382 | -103.8229
7.942 | -103.8459
7.9459 | -103.8712
7.9498 | -103.8993
7.9536 | -103.921
7.9575 | -103.9391
7.9614 | -103.9484
7.9652 | -103.9614
7.9691 | -103.9714
7.9729 | -103.9689
7.9768 | -103.9627
7.9807 | -103.9509
7.9845 | -103.9387
7.9884 | -103.9228
7.9923 | -103.9049
7.9961 | -103.8805
8 | -103.8598
8.0039 | -103.8369
8.0077 | -103.8103
8.0116 | -103.7814
8.0155 | -103.7539
8.0193 | -103.7289
8.0232 | -103.7061
8.0271 | -103.6875
8.0309 | -103.6722
8.0348 | -103.6638
8.0386 | -103.6636
8.0425 | -103.6642
8.0464 | -103.671
8.0502 | -103.6814
8.0541 | -103.6965
8.058 | -103.7139
8.0618 | -103.7367
8.0657 | -103.7539
8.0696 | -103.7674
8.0734 | -103.7853
8.0773 | -103.7951
8.0812 | -103.8032
8.085 | -103.8063
8.0889 | -103.8046
8.0928 | -103.8049
8.0966 | -103.8024
8.1005 | -103.7975
8.1043 | -103.7933
8.1082 | -103.7925
8.1121 | -103.7917
8.1159 | -103.7946
8.1198 | -103.7949
8.1237 | -103.8078
8.1275 | -103.8237
8.1314 | -103.8468
8.1353 | -103.8714
8.1391 | -103.8984
8.143 | -103.9274
8.1469 | -103.9555
8.1507 | -103.9836
8.1546 | -104.0065
8.1585 | -104.0228
8.1623 | -104.0352
8.1662 | -104.0418
8.17 | -104.0414
8.1739 | -104.0309
8.1778 | -104.0204
8.1816 | -104.012
8.1855 | -104.0115
8.1894 | -104.0112
8.1932 | -104.0056
8.1971 | -104.0136
8.201 | -104.0253
8.2048 | -104.0297
8.2087 | -104.0442
8.2126 | -104.056
8.2164 | -104.0662
8.2203 | -104.0721
8.2242 | -104.0681
8.228 | -104.0585
8.2319 | -104.0409
8.2357 | -104.0114
8.2396 | -103.9756
8.2435 | -103.9238
8.2473 | -103.8724
8.2512 | -103.8179
8.2551 | -103.7667
8.2589 | -103.7177
8.2628 | -103.6721
8.2667 | -103.6366
8.2705 | -103.6093
8.2744 | -103.5917
8.2783 | -103.5842
8.2821 | -103.5859
8.286 | -103.5911
8.2899 | -103.6042
8.2937 | -103.6211
8.2976 | -103.6325
8.3014 | -103.6465
8.3053 | -103.6645
8.3092 | -103.6816
8.313 | -103.697
8.3169 | -103.7013
8.3208 | -103.6956
8.3246 | -103.6886
8.3285 | -103.6716
8.3324 | -103.6481
8.3362 | -103.6168
8.3401 | -103.5879
8.344 | -103.5523
8.3478 | -103.5188
8.3517 | -103.4925
8.3556 | -103.4728
8.3594 | -103.4607
8.3633 | -103.4729
8.3672 | -103.4715
8.371 | -103.4729
8.3749 | -103.4812
8.3787 | -103.4896
8.3826 | -103.5018
8.3865 | -103.5148
8.3903 | -103.5257
8.3942 | -103.5342
8.3981 | -103.5418
8.4019 | -103.5442
8.4058 | -103.5465
8.4097 | -103.5483
8.4135 | -103.5436
8.4174 | -103.536
8.4213 | -103.5319
8.4251 | -103.5244
8.429 | -103.5145
8.4329 | -103.5083
8.4367 | -103.5005
8.4406 | -103.4961
8.4444 | -103.5008
8.4483 | -103.503
8.4522 | -103.5058
8.456 | -103.5132
8.4599 | -103.5236
8.4638 | -103.5327
8.4676 | -103.5481
8.4715 | -103.5666
8.4754 | -103.583
8.4792 | -103.6013
8.4831 | -103.6233
8.487 | -103.648
8.4908 | -103.6763
8.4947 | -103.7019
8.4986 | -103.7277
8.5024 | -103.7511
8.5063 | -103.7756
8.5101 | -103.7984
8.514 | -103.822
8.5179 | -103.8406
8.5217 | -103.8597
8.5256 | -103.877
8.5295 | -103.8959
8.5333 | -103.9078
8.5372 | -103.9211
8.5411 | -103.943
8.5449 | -103.9634
8.5488 | -103.9781
8.5527 | -103.9966
8.5565 | -104.0161
8.5604 | -104.0348
8.5643 | -104.0567
8.5681 | -104.0756
8.572 | -104.0874
8.5758 | -104.1036
8.5797 | -104.1185
8.5836 | -104.1311
8.5874 | -104.1432
8.5913 | -104.1505
8.5952 | -104.153
8.599 | -104.1474
8.6029 | -104.1417
8.6068 | -104.1311
8.6106 | -104.1093
8.6145 | -104.0811
8.6184 | -104.0565
8.6222 | -104.0327
8.6261 | -104.0018
8.63 | -103.9742
8.6338 | -103.949
8.6377 | -103.9225
8.6415 | -103.9008
8.6454 | -103.8905
8.6493 | -103.8792
8.6531 | -103.8691
8.657 | -103.8718
8.6609 | -103.8843
8.6647 | -103.895
8.6686 | -103.9081
8.6725 | -103.9298
8.6763 | -103.9539
8.6802 | -103.98
8.6841 | -104.0099
8.6879 | -104.0448
8.6918 | -104.0811
8.6957 | -104.1174
8.6995 | -104.1533
8.7034 | -104.1895
8.7072 | -104.2238
8.7111 | -104.2573
8.715 | -104.2912
8.7188 | -104.3223
8.7227 | -104.3478
8.7266 | -104.3682
8.7304 | -104.3842
8.7343 | -104.3933
8.7382 | -104.3989
8.742 | -104.3957
8.7459 | -104.3825
8.7498 | -104.3662
8.7536 | -104.3472
8.7575 | -104.3258
8.7614 | -104.3035
8.7652 | -104.2886
8.7691 | -104.2785
8.7729 | -104.2702
8.7768 | -104.2725
8.7807 | -104.2869
8.7845 | -104.303
8.7884 | -104.3228
8.7923 | -104.348
8.7961 | -104.3717
8.8 | -104.3944
8.8039 | -104.4178
8.8077 | -104.4346
8.8116 | -104.4514
8.8155 | -104.4603
8.8193 | -104.4587
8.8232 | -104.4517
8.8271 | -104.4443
8.8309 | -104.4301
8.8348 | -104.4125
8.8386 | -104.389
8.8425 | -104.355
8.8464 | -104.3222
8.8502 | -104.2921
8.8541 | -104.2591
8.858 | -104.2238
8.8618 | -104.1907
8.8657 | -104.1574
8.8696 | -104.1244
8.8734 | -104.091
8.8773 | -104.0557
8.8812 | -104.025
8.885 | -103.9929
8.8889 | -103.9656
8.8928 | -103.9361
8.8966 | -103.9043
8.9005 | -103.8736
8.9043 | -103.8464
8.9082 | -103.818
8.9121 | -103.7939
8.9159 | -103.7733
8.9198 | -103.7677
8.9237 | -103.758
8.9275 | -103.7476
8.9314 | -103.7403
8.9353 | -103.7322
8.9391 | -103.7326
8.943 | -103.7341
8.9469 | -103.7397
8.9507 | -103.7511
8.9546 | -103.7633
8.9585 | -103.7758
8.9623 | -103.7879
8.9662 | -103.8019
8.97 | -103.8166
8.9739 | -103.8328
8.9778 | -103.8446
8.9816 | -103.8551
8.9855 | -103.8718
8.9894 | -103.8919
8.9932 | -103.9114
8.9971 | -103.9239
9.001 | -103.9416
9.0048 | -103.9589
9.0087 | -103.9758
9.0126 | -103.9909
9.0164 | -104.0015
9.0203 | -104.0085
9.0242 | -104.012
9.028 | -104.017
9.0319 | -104.0262
9.0357 | -104.0305
9.0396 | -104.0343
9.0435 | -104.0395
9.0473 | -104.0505
9.0512 | -104.06
9.0551 | -104.0702
9.0589 | -104.0863
9.0628 | -104.1045
9.0667 | -104.1245
9.0705 | -104.144
9.0744 | -104.1578
9.0783 | -104.1714
9.0821 | -104.1849
9.086 | -104.1989
9.0899 | -104.2075
9.0937 | -104.2118
9.0976 | -104.2187
9.1014 | -104.228
9.1053 | -104.237
9.1092 | -104.2412
9.113 | -104.2474
9.1169 | -104.2497
9.1208 | -104.2443
9.1246 | -104.2338
9.1285 | -104.2266
9.1324 | -104.2262
9.1362 | -104.219
9.1401 | -104.2114
9.144 | -104.2059
9.1478 | -104.1983
9.1517 | -104.1917
9.1556 | -104.1842
9.1594 | -104.1791
9.1633 | -104.1738
9.1671 | -104.1682
9.171 | -104.1632
9.1749 | -104.1592
9.1787 | -104.1536
9.1826 | -104.1511
9.1865 | -104.1546
9.1903 | -104.1555
9.1942 | -104.1544
9.1981 | -104.1562
9.2019 | -104.1547
9.2058 | -104.1505
9.2097 | -104.1442
9.2135 | -104.1415
9.2174 | -104.1383
9.2213 | -104.1318
9.2251 | -104.1304
9.229 | -104.1301
9.2328 | -104.1255
9.2367 | -104.1235
9.2406 | -104.1271
9.2444 | -104.1389
9.2483 | -104.1546
9.2522 | -104.1766
9.256 | -104.2032
9.2599 | -104.2365
9.2638 | -104.2716
9.2676 | -104.3095
9.2715 | -104.349
9.2754 | -104.3877
9.2792 | -104.428
9.2831 | -104.4669
9.287 | -104.5052
9.2908 | -104.5422
9.2947 | -104.5738
9.2986 | -104.5978
9.3024 | -104.62
9.3063 | -104.6427
9.3101 | -104.6586
9.314 | -104.6692
9.3179 | -104.6721
9.3217 | -104.67
9.3256 | -104.6686
9.3295 | -104.6672
9.3333 | -104.6651
9.3372 | -104.6648
9.3411 | -104.6697
9.3449 | -104.6716
9.3488 | -104.6715
9.3527 | -104.673
9.3565 | -104.6723
9.3604 | -104.6694
9.3643 | -104.6692
9.3681 | -104.6714
9.372 | -104.674
9.3758 | -104.6722
9.3797 | -104.6703
9.3836 | -104.6702
9.3874 | -104.6726
9.3913 | -104.6711
9.3952 | -104.6659
9.399 | -104.6605
9.4029 | -104.6537
9.4068 | -104.6476
9.4106 | -104.6438
9.4145 | -104.6385
9.4184 | -104.6375
9.4222 | -104.6407
9.4261 | -104.6464
9.43 | -104.6586
9.4338 | -104.6758
9.4377 | -104.6906
9.4415 | -104.7066
9.4454 | -104.7246
9.4493 | -104.7451
9.4531 | -104.769
9.457 | -104.7861
9.4609 | -104.7903
9.4647 | -104.7869
9.4686 | -104.7747
9.4725 | -104.7515
9.4763 | -104.6987
9.4802 | -104.6625
9.4841 | -104.6235
9.4879 | -104.5817
9.4918 | -104.5378
9.4957 | -104.4958
9.4995 | -104.4534
9.5034 | -104.4144
9.5072 | -104.38
9.5111 | -104.3454
9.515 | -104.3207
9.5188 | -104.3008
9.5227 | -104.2836
9.5266 | -104.2735
9.5304 | -104.2735
9.5343 | -104.2822
9.5382 | -104.2936
9.542 | -104.3062
9.5459 | -104.3254
9.5498 | -104.3458
9.5536 | -104.364
9.5575 | -104.3822
9.5614 | -104.3991
9.5652 | -104.4125
9.5691 | -104.4228
9.5729 | -104.4366
9.5768 | -104.4486
9.5807 | -104.4537
9.5845 | -104.4607
9.5884 | -104.4663
9.5923 | -104.4676
9.5961 | -104.4739
9.6 | -104.4815
9.6039 | -104.4866
9.6077 | -104.489
9.6116 | -104.4958
9.6155 | -104.5036
9.6193 | -104.5125
9.6232 | -104.5239
9.6271 | -104.5356
9.6309 | -104.5491
9.6348 | -104.5671
9.6386 | -104.5847
9.6425 | -104.6023
9.6464 | -104.6181
9.6502 | -104.6363
9.6541 | -104.6612
9.658 | -104.6849
9.6618 | -104.7069
9.6657 | -104.7265
9.6696 | -104.7423
9.6734 | -104.7577
9.6773 | -104.768
9.6812 | -104.7742
9.685 | -104.7831
9.6889 | -104.7875
9.6928 | -104.7859
9.6966 | -104.7808
9.7005 | -104.7741
9.7043 | -104.7633
9.7082 | -104.7541
9.7121 | -104.7424
9.7159 | -104.7256
9.7198 | -104.7137
9.7237 | -104.708
9.7275 | -104.6994
9.7314 | -104.6894
9.7353 | -104.6816
9.7391 | -104.6762
9.743 | -104.6823
9.7469 | -104.6903
9.7507 | -104.7034
9.7546 | -104.7112
9.7585 | -104.7195
9.7623 | -104.7338
9.7662 | -104.7481
9.77 | -104.7624
9.7739 | -104.7714
9.7778 | -104.7743
9.7816 | -104.7782
9.7855 | -104.7777
9.7894 | -104.7708
9.7932 | -104.757
9.7971 | -104.7394
9.801 | -104.7215
9.8048 | -104.7051
9.8087 | -104.689
9.8126 | -104.6687
9.8164 | -104.6479
9.8203 | -104.6289
9.8242 | -104.6165
9.828 | -104.6032
9.8319 | -104.5911
9.8357 | -104.5903
9.8396 | -104.5952
9.8435 | -104.5991
9.8473 | -104.6044
9.8512 | -104.6091
9.8551 | -104.6157
9.8589 | -104.6224
9.8628 | -104.6284
9.8667 | -104.6342
9.8705 | -104.6406
9.8744 | -104.6466
9.8783 | -104.6555
9.8821 | -104.6666
9.886 | -104.6739
9.8899 | -104.6814
9.8937 | -104.6912
9.8976 | -104.7064
9.9014 | -104.7223
9.9053 | -104.7356
9.9092 | -104.7519
9.913 | -104.7666
9.9169 | -104.7838
9.9208 | -104.8059
9.9246 | -104.8236
9.9285 | -104.8398
9.9324 | -104.8528
9.9362 | -104.8658
9.9401 | -104.8738
9.944 | -104.8784
9.9478 | -104.8804
9.9517 | -104.8793
9.9556 | -104.8808
9.9594 | -104.8773
9.9633 | -104.8673
9.9672 | -104.8556
9.971 | -104.8407
9.9749 | -104.8253
9.9787 | -104.8066
9.9826 | -104.7912
9.9865 | -104.7772
9.9903 | -104.7649
9.9942 | -104.7531
9.9981 | -104.7428
10.0019 | -104.734
10.0058 | -104.7222
10.0097 | -104.7101
10.0135 | -104.6994
10.0174 | -104.6933
10.0213 | -104.6855
10.0251 | -104.6781
10.029 | -104.6714
10.0328 | -104.6779
10.0367 | -104.68
10.0406 | -104.6785
10.0444 | -104.6758
10.0483 | -104.6754
10.0522 | -104.6779
10.056 | -104.6834
10.0599 | -104.6923
10.0638 | -104.7
10.0676 | -104.7061
10.0715 | -104.7148
10.0754 | -104.7199
10.0792 | -104.7236
10.0831 | -104.7288
10.087 | -104.7325
10.0908 | -104.733
10.0947 | -104.7299
10.0986 | -104.7294
10.1024 | -104.7271
10.1063 | -104.7259
10.1101 | -104.725
10.114 | -104.7212
10.1179 | -104.7189
10.1217 | -104.7151
10.1256 | -104.7116
10.1295 | -104.7057
10.1333 | -104.705
10.1372 | -104.7079
10.1411 | -104.7152
10.1449 | -104.7242
10.1488 | -104.7315
10.1527 | -104.7407
10.1565 | -104.7549
10.1604 | -104.77
10.1643 | -104.7825
10.1681 | -104.7958
10.172 | -104.808
10.1758 | -104.8221
10.1797 | -104.8355
10.1836 | -104.8498
10.1874 | -104.8689
10.1913 | -104.8893
10.1952 | -104.9111
10.199 | -104.9291
10.2029 | -104.9462
10.2068 | -104.9587
10.2106 | -104.9717
10.2145 | -104.9874
10.2184 | -104.9993
10.2222 | -105.0071
10.2261 | -105.0192
10.23 | -105.0312
10.2338 | -105.035
10.2377 | -105.0358
10.2415 | -105.0344
10.2454 | -105.0304
10.2493 | -105.0206
10.2531 | -105.0086
10.257 | -104.9929
10.2609 | -104.9784
10.2647 | -104.9592
10.2686 | -104.9374
10.2725 | -104.9179
10.2763 | -104.8943
10.2802 | -104.8738
10.2841 | -104.8464
10.2879 | -104.8156
10.2918 | -104.7863
10.2957 | -104.761
10.2995 | -104.7418
10.3034 | -104.7241
10.3072 | -104.7038
10.3111 | -104.6842
10.315 | -104.6702
10.3188 | -104.6596
10.3227 | -104.6525
10.3266 | -104.6465
10.3304 | -104.6415
10.3343 | -104.6391
10.3382 | -104.6387
10.342 | -104.6412
10.3459 | -104.6447
10.3498 | -104.6525
10.3536 | -104.6599
10.3575 | -104.67
10.3614 | -104.6811
10.3652 | -104.6895
10.3691 | -104.7009
10.3729 | -104.7146
10.3768 | -104.7229
10.3807 | -104.7305
10.3845 | -104.7412
10.3884 | -104.7546
10.3923 | -104.7646
10.3961 | -104.7753
10.4 | -104.7852
10.4039 | -104.7909
10.4077 | -104.7962
10.4116 | -104.7988
10.4155 | -104.7999
10.4193 | -104.8002
10.4232 | -104.7923
10.4271 | -104.7876
10.4309 | -104.7785
10.4348 | -104.771
10.4386 | -104.7649
10.4425 | -104.7541
10.4464 | -104.7409
10.4502 | -104.7305
10.4541 | -104.7205
10.458 | -104.7116
10.4618 | -104.7078
10.4657 | -104.7064
10.4696 | -104.7049
10.4734 | -104.7038
10.4773 | -104.7032
10.4812 | -104.708
10.485 | -104.7157
10.4889 | -104.7259
10.4928 | -104.7387
10.4966 | -104.75
10.5005 | -104.7588
10.5043 | -104.7719
10.5082 | -104.7856
10.5121 | -104.8013
10.5159 | -104.8197
10.5198 | -104.8378
10.5237 | -104.8564
10.5275 | -104.877
10.5314 | -104.896
10.5353 | -104.9138
10.5391 | -104.927
10.543 | -104.9428
10.5469 | -104.9568
10.5507 | -104.9678
10.5546 | -104.9802
10.5585 | -104.9953
10.5623 | -105.0102
10.5662 | -105.0244
10.57 | -105.038
10.5739 | -105.0542
10.5778 | -105.0679
10.5816 | -105.0826
10.5855 | -105.0963
10.5894 | -105.1021
10.5932 | -105.0976
10.5971 | -105.1012
10.601 | -105.101
10.6048 | -105.0923
10.6087 | -105.0879
10.6126 | -105.0842
10.6164 | -105.0724
10.6203 | -105.0571
10.6242 | -105.0443
10.628 | -105.031
10.6319 | -105.0139
10.6357 | -104.9962
10.6396 | -104.9775
10.6435 | -104.9596
10.6473 | -104.945
10.6512 | -104.9322
10.6551 | -104.9237
10.6589 | -104.9165
10.6628 | -104.9153
10.6667 | -104.9108
10.6705 | -104.905
10.6744 | -104.903
10.6783 | -104.9041
10.6821 | -104.9079
10.686 | -104.9121
10.6899 | -104.9181
10.6937 | -104.9279
10.6976 | -104.9343
10.7014 | -104.9406
10.7053 | -104.9458
10.7092 | -104.9497
10.713 | -104.9507
10.7169 | -104.9511
10.7208 | -104.9525
10.7246 | -104.9534
10.7285 | -104.9544
10.7324 | -104.9549
10.7362 | -104.9555
10.7401 | -104.9579
10.744 | -104.9659
10.7478 | -104.9725
10.7517 | -104.9812
10.7556 | -104.9881
10.7594 | -104.9961
10.7633 | -105.0058
10.7672 | -105.0095
10.771 | -105.0091
10.7749 | -105.0139
10.7787 | -105.0227
10.7826 | -105.0273
10.7865 | -105.0321
10.7903 | -105.0403
10.7942 | -105.0433
10.7981 | -105.045
10.8019 | -105.0461
10.8058 | -105.0438
10.8097 | -105.0397
10.8135 | -105.0325
10.8174 | -105.0225
10.8213 | -105.0082
10.8251 | -104.995
10.829 | -104.9836
10.8329 | -104.9692
10.8367 | -104.9532
10.8406 | -104.9424
10.8444 | -104.9371
10.8483 | -104.9332
10.8522 | -104.9286
10.856 | -104.9308
10.8599 | -104.9374
10.8638 | -104.949
10.8676 | -104.9624
10.8715 | -104.9777
10.8754 | -104.9958
10.8792 | -105.017
10.8831 | -105.0393
10.887 | -105.0613
10.8908 | -105.0865
10.8947 | -105.1112
10.8986 | -105.1345
10.9024 | -105.1556
10.9063 | -105.1741
10.9101 | -105.1868
10.914 | -105.196
10.9179 | -105.1978
10.9217 | -105.1866
10.9256 | -105.1767
10.9295 | -105.1634
10.9333 | -105.1388
10.9372 | -105.1058
10.9411 | -105.0674
10.9449 | -105.0259
10.9488 | -104.9786
10.9527 | -104.9278
10.9565 | -104.8741
10.9604 | -104.8199
10.9643 | -104.7714
10.9681 | -104.724
10.972 | -104.6798
10.9758 | -104.6376
10.9797 | -104.6008
10.9836 | -104.5708
10.9874 | -104.5474
10.9913 | -104.5328
10.9952 | -104.528
10.999 | -104.5252
11.0029 | -104.5266
11.0068 | -104.5383
11.0106 | -104.5633
11.0145 | -104.5909
11.0184 | -104.6229
11.0222 | -104.6634
11.0261 | -104.7092
11.03 | -104.7578
11.0338 | -104.8101
11.0377 | -104.8614
11.0415 | -104.9099
11.0454 | -104.9611
11.0493 | -105.0056
11.0531 | -105.0478
11.057 | -105.0827
11.0609 | -105.1124
11.0647 | -105.1411
11.0686 | -105.1692
11.0725 | -105.1933
11.0763 | -105.2119
11.0802 | -105.2301
11.0841 | -105.2369
11.0879 | -105.242
11.0918 | -105.2461
11.0957 | -105.2455
11.0995 | -105.2461
11.1034 | -105.2418
11.1072 | -105.2384
11.1111 | -105.2362
11.115 | -105.2327
11.1188 | -105.2325
11.1227 | -105.2313
11.1266 | -105.2357
11.1304 | -105.2414
11.1343 | -105.2504
11.1382 | -105.2627
11.142 | -105.2768
11.1459 | -105.2919
11.1498 | -105.3159
11.1536 | -105.3305
11.1575 | -105.3405
11.1614 | -105.3573
11.1652 | -105.3776
11.1691 | -105.3928
11.1729 | -105.4053
11.1768 | -105.4216
11.1807 | -105.4335
11.1845 | -105.4337
11.1884 | -105.4305
11.1923 | -105.4266
11.1961 | -105.4155
11.2 | -105.4038
11.2039 | -105.3891
11.2077 | -105.3696
11.2116 | -105.3512
11.2155 | -105.331
11.2193 | -105.316
11.2232 | -105.2987
11.2271 | -105.2801
11.2309 | -105.2675
11.2348 | -105.2586
11.2386 | -105.2529
11.2425 | -105.2452
11.2464 | -105.2389
11.2502 | -105.2359
11.2541 | -105.2393
11.258 | -105.2495
11.2618 | -105.2593
11.2657 | -105.2751
11.2696 | -105.2889
11.2734 | -105.302
11.2773 | -105.3188
11.2812 | -105.3318
11.285 | -105.3431
11.2889 | -105.3529
11.2928 | -105.363
11.2966 | -105.3696
11.3005 | -105.3742
11.3043 | -105.3741
11.3082 | -105.3699
11.3121 | -105.3651
11.3159 | -105.3593
11.3198 | -105.3494
11.3237 | -105.3428
11.3275 | -105.3314
11.3314 | -105.3146
11.3353 | -105.2998
11.3391 | -105.2828
11.343 | -105.2645
11.3469 | -105.2518
11.3507 | -105.2493
11.3546 | -105.2463
11.3585 | -105.2422
11.3623 | -105.2389
11.3662 | -105.236
11.37 | -105.2334
11.3739 | -105.2287
11.3778 | -105.2246
11.3816 | -105.2197
11.3855 | -105.2105
11.3894 | -105.2029
11.3932 | -105.1909
11.3971 | -105.1741
11.401 | -105.1588
11.4048 | -105.1456
11.4087 | -105.1361
11.4126 | -105.1236
11.4164 | -105.1071
11.4203 | -105.0922
11.4242 | -105.0765
11.428 | -105.0624
11.4319 | -105.0483
11.4357 | -105.0393
11.4396 | -105.0336
11.4435 | -105.0278
11.4473 | -105.028
11.4512 | -105.0315
11.4551 | -105.0364
11.4589 | -105.0434
11.4628 | -105.0545
11.4667 | -105.069
11.4705 | -105.0858
11.4744 | -105.1036
11.4783 | -105.1181
11.4821 | -105.1333
11.486 | -105.1512
11.4899 | -105.1663
11.4937 | -105.1793
11.4976 | -105.1913
11.5014 | -105.1944
11.5053 | -105.1984
11.5092 | -105.1998
11.513 | -105.1932
11.5169 | -105.1813
11.5208 | -105.172
11.5246 | -105.16
11.5285 | -105.1454
11.5324 | -105.1294
11.5362 | -105.1144
11.5401 | -105.1023
11.544 | -105.0878
11.5478 | -105.0704
11.5517 | -105.0568
11.5556 | -105.0475
11.5594 | -105.0415
11.5633 | -105.0363
11.5671 | -105.0291
11.571 | -105.024
11.5749 | -105.0203
11.5787 | -105.021
11.5826 | -105.0249
11.5865 | -105.0277
11.5903 | -105.0366
11.5942 | -105.0517
11.5981 | -105.0679
11.6019 | -105.0805
11.6058 | -105.099
11.6097 | -105.1221
11.6135 | -105.1445
11.6174 | -105.1677
11.6213 | -105.1968
11.6251 | -105.2287
11.629 | -105.2588
11.6328 | -105.2888
11.6367 | -105.3188
11.6406 | -105.3475
11.6444 | -105.3766
11.6483 | -105.4042
11.6522 | -105.4282
11.656 | -105.4475
11.6599 | -105.4681
11.6638 | -105.4893
11.6676 | -105.5119
11.6715 | -105.5286
11.6754 | -105.5437
11.6792 | -105.5579
11.6831 | -105.5679
11.687 | -105.574
11.6908 | -105.5803
11.6947 | -105.5898
11.6986 | -105.5951
11.7024 | -105.5945
11.7063 | -105.5842
11.7101 | -105.5771
11.714 | -105.5665
11.7179 | -105.5594
11.7217 | -105.5477
11.7256 | -105.5279
11.7295 | -105.5046
11.7333 | -105.4834
11.7372 | -105.4612
11.7411 | -105.4375
11.7449 | -105.4144
11.7488 | -105.3906
11.7527 | -105.3639
11.7565 | -105.3374
11.7604 | -105.3101
11.7643 | -105.2871
11.7681 | -105.2666
11.772 | -105.2519
11.7758 | -105.235
11.7797 | -105.2182
11.7836 | -105.206
11.7874 | -105.1952
11.7913 | -105.1844
11.7952 | -105.1741
11.799 | -105.1683
11.8029 | -105.1639
11.8068 | -105.1599
11.8106 | -105.1562
11.8145 | -105.1507
11.8184 | -105.1472
11.8222 | -105.147
11.8261 | -105.1485
11.83 | -105.1491
11.8338 | -105.1492
11.8377 | -105.1473
11.8415 | -105.1438
11.8454 | -105.1377
11.8493 | -105.1362
11.8531 | -105.138
11.857 | -105.1372
11.8609 | -105.1356
11.8647 | -105.1298
11.8686 | -105.126
11.8725 | -105.125
11.8763 | -105.1218
11.8802 | -105.1163
11.8841 | -105.1112
11.8879 | -105.1117
11.8918 | -105.1119
11.8957 | -105.1082
11.8995 | -105.1075
11.9034 | -105.1115
11.9072 | -105.1124
11.9111 | -105.1118
11.915 | -105.1123
11.9188 | -105.1134
11.9227 | -105.1145
11.9266 | -105.1153
11.9304 | -105.1133
11.9343 | -105.1142
11.9382 | -105.1108
11.942 | -105.1059
11.9459 | -105.1038
11.9498 | -105.1002
11.9536 | -105.0952
11.9575 | -105.0879
11.9614 | -105.0821
11.9652 | -105.0752
11.9691 | -105.0673
11.9729 | -105.0573
11.9768 | -105.0456
11.9807 | -105.0364
11.9845 | -105.0216
11.9884 | -105.0078
11.9923 | -104.9945
11.9961 | -104.983
12 | -104.974
12.0039 | -104.9654
12.0077 | -104.9612
12.0116 | -104.9627
12.0155 | -104.9655
12.0193 | -104.9732
12.0232 | -104.9797
12.0271 | -104.9882
12.0309 | -105
12.0348 | -105.0178
12.0386 | -105.0431
12.0425 | -105.0648
12.0464 | -105.087
12.0502 | -105.1117
12.0541 | -105.1398
12.058 | -105.1679
12.0618 | -105.1918
12.0657 | -105.2169
12.0696 | -105.2433
12.0734 | -105.2616
12.0773 | -105.2732
12.0812 | -105.2854
12.085 | -105.2972
12.0889 | -105.3044
12.0928 | -105.309
12.0966 | -105.3096
12.1005 | -105.3109
12.1043 | -105.3117
12.1082 | -105.3068
12.1121 | -105.2992
12.1159 | -105.2902
12.1198 | -105.2759
12.1237 | -105.2622
12.1275 | -105.251
12.1314 | -105.2384
12.1353 | -105.2245
12.1391 | -105.2068
12.143 | -105.1887
12.1469 | -105.1771
12.1507 | -105.1674
12.1546 | -105.1556
12.1585 | -105.1496
12.1623 | -105.1404
12.1662 | -105.1283
12.17 | -105.1154
12.1739 | -105.1028
12.1778 | -105.0915
12.1816 | -105.0818
12.1855 | -105.0709
12.1894 | -105.0652
12.1932 | -105.0589
12.1971 | -105.0541
12.201 | -105.0476
12.2048 | -105.0405
12.2087 | -105.0336
12.2126 | -105.028
12.2164 | -105.0311
12.2203 | -105.0368
12.2242 | -105.0423
12.228 | -105.0492
12.2319 | -105.06
12.2357 | -105.0751
12.2396 | -105.0902
12.2435 | -105.1106
12.2473 | -105.1372
12.2512 | -105.1648
12.2551 | -105.1931
12.2589 | -105.2201
12.2628 | -105.2531
12.2667 | -105.2799
12.2705 | -105.3058
12.2744 | -105.3342
12.2783 | -105.3611
12.2821 | -105.3883
12.286 | -105.4144
12.2899 | -105.4408
12.2937 | -105.4654
12.2976 | -105.4871
12.3014 | -105.504
12.3053 | -105.5178
12.3092 | -105.5318
12.313 | -105.5458
12.3169 | -105.5568
12.3208 | -105.5635
12.3246 | -105.5611
12.3285 | -105.5556
12.3324 | -105.5475
12.3362 | -105.5333
12.3401 | -105.5196
12.344 | -105.5038
12.3478 | -105.484
12.3517 | -105.4617
12.3556 | -105.4393
12.3594 | -105.413
12.3633 | -105.3894
12.3672 | -105.3614
12.371 | -105.3337
12.3749 | -105.3098
12.3787 | -105.2874
12.3826 | -105.2674
12.3865 | -105.248
12.3903 | -105.2261
12.3942 | -105.2046
12.3981 | -105.1846
12.4019 | -105.1732
12.4058 | -105.1652
12.4097 | -105.1562
12.4135 | -105.1444
12.4174 | -105.1349
12.4213 | -105.1288
12.4251 | -105.1233
12.429 | -105.1188
12.4329 | -105.1139
12.4367 | -105.1133
12.4406 | -105.1075
12.4444 | -105.1006
12.4483 | -105.0955
12.4522 | -105.0903
12.456 | -105.084
12.4599 | -105.0768
12.4638 | -105.0653
12.4676 | -105.0545
12.4715 | -105.0433
12.4754 | -105.0305
12.4792 | -105.0203
12.4831 | -105.0138
12.487 | -105.0057
12.4908 | -104.9958
12.4947 | -104.9857
12.4986 | -104.9761
12.5024 | -104.9688
12.5063 | -104.9648
12.5101 | -104.9601
12.514 | -104.9585
12.5179 | -104.9608
12.5217 | -104.9617
12.5256 | -104.9662
12.5295 | -104.9703
12.5333 | -104.9763
12.5372 | -104.9833
12.5411 | -104.9924
12.5449 | -105.0007
12.5488 | -105.0108
12.5527 | -105.0206
12.5565 | -105.0296
12.5604 | -105.043
12.5643 | -105.0572
12.5681 | -105.0701
12.572 | -105.085
12.5758 | -105.1004
12.5797 | -105.1153
12.5836 | -105.1295
12.5874 | -105.145
12.5913 | -105.1625
12.5952 | -105.1839
12.599 | -105.2018
12.6029 | -105.2136
12.6068 | -105.228
12.6106 | -105.2426
12.6145 | -105.2554
12.6184 | -105.2654
12.6222 | -105.2787
12.6261 | -105.2934
12.63 | -105.3042
12.6338 | -105.3149
12.6377 | -105.3236
12.6415 | -105.3336
12.6454 | -105.3468
12.6493 | -105.3573
12.6531 | -105.3694
12.657 | -105.3853
12.6609 | -105.396
12.6647 | -105.4058
12.6686 | -105.4205
12.6725 | -105.4363
12.6763 | -105.4544
12.6802 | -105.4737
12.6841 | -105.4905
12.6879 | -105.5065
12.6918 | -105.5204
12.6957 | -105.5363
12.6995 | -105.5501
12.7034 | -105.5644
12.7072 | -105.5718
12.7111 | -105.579
12.715 | -105.5839
12.7188 | -105.5854
12.7227 | -105.5893
12.7266 | -105.5897
12.7304 | -105.5876
12.7343 | -105.5839
12.7382 | -105.5794
12.742 | -105.5713
12.7459 | -105.5591
12.7498 | -105.5432
12.7536 | -105.5268
12.7575 | -105.5086
12.7614 | -105.4907
12.7652 | -105.4687
12.7691 | -105.4468
12.7729 | -105.4275
12.7768 | -105.4062
12.7807 | -105.3905
12.7845 | -105.3788
12.7884 | -105.364
12.7923 | -105.347
12.7961 | -105.3293
12.8 | -105.3109
12.8039 | -105.291
12.8077 | -105.27
12.8116 | -105.2498
12.8155 | -105.2267
12.8193 | -105.2079
12.8232 | -105.191
12.8271 | -105.177
12.8309 | -105.164
12.8348 | -105.1534
12.8386 | -105.1426
12.8425 | -105.1306
12.8464 | -105.1221
12.8502 | -105.1107
12.8541 | -105.1015
12.858 | -105.0995
12.8618 | -105.1011
12.8657 | -105.1032
12.8696 | -105.1073
12.8734 | -105.11
12.8773 | -105.1169
12.8812 | -105.1234
12.885 | -105.1338
12.8889 | -105.1427
12.8928 | -105.1534
12.8966 | -105.1676
12.9005 | -105.184
12.9043 | -105.1971
12.9082 | -105.2076
12.9121 | -105.2226
12.9159 | -105.2391
12.9198 | -105.2578
12.9237 | -105.2754
12.9275 | -105.2911
12.9314 | -105.3108
12.9353 | -105.3327
12.9391 | -105.3505
12.943 | -105.3689
12.9469 | -105.3865
12.9507 | -105.403
12.9546 | -105.4181
12.9585 | -105.4306
12.9623 | -105.444
12.9662 | -105.4523
12.97 | -105.4574
12.9739 | -105.4624
12.9778 | -105.47
12.9816 | -105.4786
12.9855 | -105.4848
12.9894 | -105.4926
12.9932 | -105.4992
12.9971 | -105.5052
13.001 | -105.5094
13.0048 | -105.5111
13.0087 | -105.5095
13.0126 | -105.5106
13.0164 | -105.508
13.0203 | -105.5049
13.0242 | -105.5025
13.028 | -105.5018
13.0319 | -105.5043
13.0357 | -105.5039
13.0396 | -105.5042
13.0435 | -105.5058
13.0473 | -105.5074
13.0512 | -105.5065
13.0551 | -105.5091
13.0589 | -105.5137
13.0628 | -105.5186
13.0667 | -105.5239
13.0705 | -105.5285
13.0744 | -105.5321
13.0783 | -105.5378
13.0821 | -105.5458
13.086 | -105.5492
13.0899 | -105.5513
13.0937 | -105.5568
13.0976 | -105.5609
13.1014 | -105.5608
13.1053 | -105.5609
13.1092 | -105.5575
13.113 | -105.5533
13.1169 | -105.5542
13.1208 | -105.5555
13.1246 | -105.5557
13.1285 | -105.5588
13.1324 | -105.5577
13.1362 | -105.5563
13.1401 | -105.5561
13.144 | -105.5586
13.1478 | -105.5604
13.1517 | -105.5618
13.1556 | -105.5631
13.1594 | -105.5656
13.1633 | -105.5689
13.1671 | -105.5653
13.171 | -105.557
13.1749 | -105.5497
13.1787 | -105.5388
13.1826 | -105.5244
13.1865 | -105.5072
13.1903 | -105.4918
13.1942 | -105.4728
13.1981 | -105.4514
13.2019 | -105.4199
13.2058 | -105.3887
13.2097 | -105.3581
13.2135 | -105.3252
13.2174 | -105.2896
13.2213 | -105.2539
13.2251 | -105.2189
13.229 | -105.1871
13.2328 | -105.1496
13.2367 | -105.1089
13.2406 | -105.0728
13.2444 | -105.0416
13.2483 | -105.0115
13.2522 | -104.9848
13.256 | -104.9614
13.2599 | -104.9416
13.2638 | -104.9223
13.2676 | -104.9058
13.2715 | -104.8957
13.2754 | -104.8923
13.2792 | -104.8932
13.2831 | -104.8979
13.287 | -104.907
13.2908 | -104.9151
13.2947 | -104.9253
13.2986 | -104.9398
13.3024 | -104.9565
13.3063 | -104.9707
13.3101 | -104.9839
13.314 | -104.9998
13.3179 | -105.0148
13.3217 | -105.0322
13.3256 | -105.0491
13.3295 | -105.0618
13.3333 | -105.069
13.3372 | -105.0771
13.3411 | -105.0847
13.3449 | -105.0898
13.3488 | -105.0907
13.3527 | -105.086
13.3565 | -105.0832
13.3604 | -105.0792
13.3643 | -105.0713
13.3681 | -105.0605
13.372 | -105.0511
13.3758 | -105.0398
13.3797 | -105.0247
13.3836 | -105.0117
13.3874 | -104.9994
13.3913 | -104.9877
13.3952 | -104.9786
13.399 | -104.963
13.4029 | -104.9459
13.4068 | -104.9304
13.4106 | -104.9168
13.4145 | -104.9034
13.4184 | -104.8943
13.4222 | -104.8775
13.4261 | -104.86
13.43 | -104.8476
13.4338 | -104.8347
13.4377 | -104.8216
13.4415 | -104.8094
13.4454 | -104.7996
13.4493 | -104.7903
13.4531 | -104.7777
13.457 | -104.768
13.4609 | -104.76
13.4647 | -104.7494
13.4686 | -104.7375
13.4725 | -104.7276
13.4763 | -104.7183
13.4802 | -104.7092
13.4841 | -104.7047
13.4879 | -104.7006
13.4918 | -104.6956
13.4957 | -104.6923
13.4995 | -104.6943
13.5034 | -104.6992
13.5072 | -104.7031
13.5111 | -104.709
13.515 | -104.7154
13.5188 | -104.7213
13.5227 | -104.7335
13.5266 | -104.7503
13.5304 | -104.7683
13.5343 | -104.7871
13.5382 | -104.8076
13.542 | -104.8279
13.5459 | -104.8526
13.5498 | -104.8784
13.5536 | -104.9034
13.5575 | -104.9262
13.5614 | -104.9477
13.5652 | -104.9679
13.5691 | -104.9888
13.5729 | -105.0086
13.5768 | -105.0272
13.5807 | -105.0454
13.5845 | -105.0612
13.5884 | -105.0769
13.5923 | -105.0936
13.5961 | -105.111
13.6 | -105.1243
13.6039 | -105.133
13.6077 | -105.1404
13.6116 | -105.1415
13.6155 | -105.1346
13.6193 | -105.1321
13.6232 | -105.1298
13.6271 | -105.1209
13.6309 | -105.1087
13.6348 | -105.094
13.6386 | -105.0773
13.6425 | -105.0552
13.6464 | -105.0322
13.6502 | -105.0116
13.6541 | -104.9913
13.658 | -104.9681
13.6618 | -104.9454
13.6657 | -104.9254
13.6696 | -104.9029
13.6734 | -104.8827
13.6773 | -104.8661
13.6812 | -104.8548
13.685 | -104.8444
13.6889 | -104.8366
13.6928 | -104.8296
13.6966 | -104.8226
13.7005 | -104.8193
13.7043 | -104.8209
13.7082 | -104.827
13.7121 | -104.8369
13.7159 | -104.8482
13.7198 | -104.8578
13.7237 | -104.8701
13.7275 | -104.8843
13.7314 | -104.8989
13.7353 | -104.9165
13.7391 | -104.9384
13.743 | -104.9591
13.7469 | -104.9789
13.7507 | -105.0035
13.7546 | -105.0283
13.7585 | -105.0516
13.7623 | -105.0733
13.7662 | -105.0922
13.77 | -105.1116
13.7739 | -105.1306
13.7778 | -105.148
13.7816 | -105.1624
13.7855 | -105.1753
13.7894 | -105.1851
13.7932 | -105.1922
13.7971 | -105.1965
13.801 | -105.1975
13.8048 | -105.1965
13.8087 | -105.1959
13.8126 | -105.1941
13.8164 | -105.1907
13.8203 | -105.1812
13.8242 | -105.1716
13.828 | -105.1618
13.8319 | -105.1511
13.8357 | -105.1397
13.8396 | -105.1323
13.8435 | -105.122
13.8473 | -105.1132
13.8512 | -105.1048
13.8551 | -105.0977
13.8589 | -105.0916
13.8628 | -105.0849
13.8667 | -105.0803
13.8705 | -105.0782
13.8744 | -105.0758
13.8783 | -105.0706
13.8821 | -105.0672
13.886 | -105.0661
13.8899 | -105.0669
13.8937 | -105.0696
13.8976 | -105.0686
13.9014 | -105.067
13.9053 | -105.066
13.9092 | -105.0657
13.913 | -105.0674
13.9169 | -105.0728
13.9208 | -105.0766
13.9246 | -105.0793
13.9285 | -105.0864
13.9324 | -105.0937
13.9362 | -105.093
13.9401 | -105.0959
13.944 | -105.0982
13.9478 | -105.1005
13.9517 | -105.0986
13.9556 | -105.0992
13.9594 | -105.0966
13.9633 | -105.0903
13.9672 | -105.085
13.971 | -105.076
13.9749 | -105.0676
13.9787 | -105.0602
13.9826 | -105.055
13.9865 | -105.0481
13.9903 | -105.0404
13.9942 | -105.0311
13.9981 | -105.0225
14.0019 | -105.0153
14.0058 | -105.0059
14.0097 | -104.9962
14.0135 | -104.99
14.0174 | -104.9831
14.0213 | -104.9792
14.0251 | -104.9742
14.029 | -104.9684
14.0328 | -104.965
14.0367 | -104.9649
14.0406 | -104.964
14.0444 | -104.966
14.0483 | -104.9706
14.0522 | -104.97
14.056 | -104.9727
14.0599 | -104.9789
14.0638 | -104.9903
14.0676 | -105.0042
14.0715 | -105.0151
14.0754 | -105.0271
14.0792 | -105.0432
14.0831 | -105.0603
14.087 | -105.0757
14.0908 | -105.0922
14.0947 | -105.1048
14.0986 | -105.1176
14.1024 | -105.1318
14.1063 | -105.1431
14.1101 | -105.1531
14.114 | -105.1615
14.1179 | -105.1651
14.1217 | -105.1665
14.1256 | -105.1668
14.1295 | -105.1641
14.1333 | -105.159
14.1372 | -105.1568
14.1411 | -105.1531
14.1449 | -105.147
14.1488 | -105.1387
14.1527 | -105.1283
14.1565 | -105.1191
14.1604 | -105.1085
14.1643 | -105.1022
14.1681 | -105.0968
14.172 | -105.089
14.1758 | -105.0848
14.1797 | -105.0759
14.1836 | -105.066
14.1874 | -105.0573
14.1913 | -105.0489
14.1952 | -105.0448
14.199 | -105.0454
14.2029 | -105.0457
14.2068 | -105.0481
14.2106 | -105.0495
14.2145 | -105.0515
14.2184 | -105.0524
14.2222 | -105.0538
14.2261 | -105.0542
14.23 | -105.0576
14.2338 | -105.0622
14.2377 | -105.0668
14.2415 | -105.0691
14.2454 | -105.0703
14.2493 | -105.0708
14.2531 | -105.0667
14.257 | -105.066
14.2609 | -105.0663
14.2647 | -105.0614
14.2686 | -105.0552
14.2725 | -105.0482
14.2763 | -105.0423
14.2802 | -105.0352
14.2841 | -105.0262
14.2879 | -105.0173
14.2918 | -105.0114
14.2957 | -105.0026
14.2995 | -104.9978
14.3034 | -104.9944
14.3072 | -104.9873
14.3111 | -104.982
14.315 | -104.9769
14.3188 | -104.9716
14.3227 | -104.9681
14.3266 | -104.9664
14.3304 | -104.9651
14.3343 | -104.9651
14.3382 | -104.9669
14.342 | -104.9684
14.3459 | -104.9731
14.3498 | -104.9803
14.3536 | -104.9892
14.3575 | -104.9997
14.3614 | -105.0101
14.3652 | -105.0195
14.3691 | -105.027
14.3729 | -105.0366
14.3768 | -105.0497
14.3807 | -105.0606
14.3845 | -105.0738
14.3884 | -105.087
14.3923 | -105.0999
14.3961 | -105.1142
14.4 | -105.1255
14.4039 | -105.1344
14.4077 | -105.1445
14.4116 | -105.1521
14.4155 | -105.1605
14.4193 | -105.1707
14.4232 | -105.1805
14.4271 | -105.1866
14.4309 | -105.1899
14.4348 | -105.1921
14.4386 | -105.195
14.4425 | -105.1947
14.4464 | -105.1937
14.4502 | -105.1923
14.4541 | -105.1889
14.458 | -105.1858
14.4618 | -105.1845
14.4657 | -105.1756
14.4696 | -105.1658
14.4734 | -105.1567
14.4773 | -105.1498
14.4812 | -105.1441
14.485 | -105.1381
14.4889 | -105.1314
14.4928 | -105.123
14.4966 | -105.1154
14.5005 | -105.1075
14.5043 | -105.101
14.5082 | -105.0939
14.5121 | -105.0856
14.5159 | -105.0813
14.5198 | -105.0779
14.5237 | -105.072
14.5275 | -105.0683
14.5314 | -105.064
14.5353 | -105.0582
14.5391 | -105.0559
14.543 | -105.0543
14.5469 | -105.0532
14.5507 | -105.0519
14.5546 | -105.0516
14.5585 | -105.05
14.5623 | -105.0508
14.5662 | -105.0515
14.57 | -105.0543
14.5739 | -105.0578
14.5778 | -105.0633
14.5816 | -105.0678
14.5855 | -105.0716
14.5894 | -105.0775
14.5932 | -105.0832
14.5971 | -105.0895
14.601 | -105.0964
14.6048 | -105.1054
14.6087 | -105.1141
14.6126 | -105.1221
14.6164 | -105.128
14.6203 | -105.135
14.6242 | -105.145
14.628 | -105.1535
14.6319 | -105.1612
14.6357 | -105.1716
14.6396 | -105.1837
14.6435 | -105.1906
14.6473 | -105.197
14.6512 | -105.2028
14.6551 | -105.2102
14.6589 | -105.2156
14.6628 | -105.2179
14.6667 | -105.2189
14.6705 | -105.2189
14.6744 | -105.2173
14.6783 | -105.2151
14.6821 | -105.2107
14.686 | -105.2006
14.6899 | -105.1892
14.6937 | -105.1801
14.6976 | -105.1706
14.7014 | -105.1629
14.7053 | -105.1546
14.7092 | -105.1452
14.713 | -105.1339
14.7169 | -105.1225
14.7208 | -105.1121
14.7246 | -105.1033
14.7285 | -105.0946
14.7324 | -105.085
14.7362 | -105.0777
14.7401 | -105.0723
14.744 | -105.0686
14.7478 | -105.0687
14.7517 | -105.0702
14.7556 | -105.0695
14.7594 | -105.0685
14.7633 | -105.0712
14.7672 | -105.0767
14.771 | -105.0814
14.7749 | -105.0848
14.7787 | -105.0895
14.7826 | -105.0937
14.7865 | -105.1002
14.7903 | -105.1063
14.7942 | -105.1109
14.7981 | -105.1139
14.8019 | -105.1187
14.8058 | -105.1266
14.8097 | -105.1353
14.8135 | -105.14
14.8174 | -105.1452
14.8213 | -105.1546
14.8251 | -105.1642
14.829 | -105.1713
14.8329 | -105.1791
14.8367 | -105.187
14.8406 | -105.1942
14.8444 | -105.2004
14.8483 | -105.205
14.8522 | -105.2094
14.856 | -105.2128
14.8599 | -105.2165
14.8638 | -105.2204
14.8676 | -105.2243
14.8715 | -105.2271
14.8754 | -105.2316
14.8792 | -105.232
14.8831 | -105.2269
14.887 | -105.2234
14.8908 | -105.2185
14.8947 | -105.2096
14.8986 | -105.2041
14.9024 | -105.1984
14.9063 | -105.1895
14.9101 | -105.1816
14.914 | -105.1717
14.9179 | -105.1648
14.9217 | -105.1572
14.9256 | -105.1477
14.9295 | -105.1411
14.9333 | -105.1323
14.9372 | -105.1237
14.9411 | -105.1152
14.9449 | -105.1093
14.9488 | -105.1057
14.9527 | -105.1011
14.9565 | -105.0968
14.9604 | -105.0887
14.9643 | -105.0797
14.9681 | -105.0758
14.972 | -105.0748
14.9758 | -105.0747
14.9797 | -105.0702
14.9836 | -105.0648
14.9874 | -105.0571
14.9913 | -105.0503
14.9952 | -105.0427
14.999 | -105.0336
15.0029 | -105.0253
15.0068 | -105.0146
15.0106 | -105.0038
15.0145 | -104.9917
15.0184 | -104.9762
15.0222 | -104.9564
15.0261 | -104.9386
15.03 | -104.9173
15.0338 | -104.8899
15.0377 | -104.8617
15.0415 | -104.8334
15.0454 | -104.801
15.0493 | -104.767
15.0531 | -104.7356
15.057 | -104.7051
15.0609 | -104.674
15.0647 | -104.6453
15.0686 | -104.6173
15.0725 | -104.5911
15.0763 | -104.5628
15.0802 | -104.5366
15.0841 | -104.5123
15.0879 | -104.4907
15.0918 | -104.4727
15.0957 | -104.4588
15.0995 | -104.4487
15.1034 | -104.441
15.1072 | -104.4359
15.1111 | -104.4328
15.115 | -104.4291
15.1188 | -104.429
15.1227 | -104.4343
15.1266 | -104.4386
15.1304 | -104.4458
15.1343 | -104.4552
15.1382 | -104.468
15.142 | -104.4851
15.1459 | -104.5009
15.1498 | -104.5178
15.1536 | -104.5339
15.1575 | -104.5511
15.1614 | -104.5712
15.1652 | -104.592
15.1691 | -104.6111
15.1729 | -104.6335
15.1768 | -104.6592
15.1807 | -104.6839
15.1845 | -104.7063
15.1884 | -104.7266
15.1923 | -104.7477
15.1961 | -104.7705
15.2 | -104.7919
15.2039 | -104.8167
15.2077 | -104.8395
15.2116 | -104.8593
15.2155 | -104.8773
15.2193 | -104.8921
15.2232 | -104.9077
15.2271 | -104.9212
15.2309 | -104.9322
15.2348 | -104.9431
15.2386 | -104.9545
15.2425 | -104.9636
15.2464 | -104.971
15.2502 | -104.9774
15.2541 | -104.9811
15.258 | -104.987
15.2618 | -104.9928
15.2657 | -104.9944
15.2696 | -104.9929
15.2734 | -104.9878
15.2773 | -104.9832
15.2812 | -104.9784
15.285 | -104.9719
15.2889 | -104.9661
15.2928 | -104.9612
15.2966 | -104.9546
15.3005 | -104.9457
15.3043 | -104.937
15.3082 | -104.9292
15.3121 | -104.9188
15.3159 | -104.9087
15.3198 | -104.8981
15.3237 | -104.8889
15.3275 | -104.8779
15.3314 | -104.8652
15.3353 | -104.8541
15.3391 | -104.8419
15.343 | -104.8295
15.3469 | -104.8162
15.3507 | -104.8026
15.3546 | -104.7881
15.3585 | -104.7726
15.3623 | -104.7592
15.3662 | -104.7462
15.37 | -104.7348
15.3739 | -104.7252
15.3778 | -104.713
15.3816 | -104.6978
15.3855 | -104.6884
15.3894 | -104.6804
15.3932 | -104.6698
15.3971 | -104.6597
15.401 | -104.6521
15.4048 | -104.6465
15.4087 | -104.6432
15.4126 | -104.6379
15.4164 | -104.6344
15.4203 | -104.6375
15.4242 | -104.6377
15.428 | -104.6365
15.4319 | -104.6366
15.4357 | -104.6396
15.4396 | -104.644
15.4435 | -104.6492
15.4473 | -104.6552
15.4512 | -104.6627
15.4551 | -104.6708
15.4589 | -104.6807
15.4628 | -104.6904
15.4667 | -104.7014
15.4705 | -104.7127
15.4744 | -104.7246
15.4783 | -104.7328
15.4821 | -104.7412
15.486 | -104.7484
15.4899 | -104.7556
15.4937 | -104.7644
15.4976 | -104.7732
15.5014 | -104.777
15.5053 | -104.7804
15.5092 | -104.7824
15.513 | -104.7857
15.5169 | -104.788
15.5208 | -104.7861
15.5246 | -104.7803
15.5285 | -104.7748
15.5324 | -104.7675
15.5362 | -104.7552
15.5401 | -104.7414
15.544 | -104.7251
15.5478 | -104.7074
15.5517 | -104.6872
15.5556 | -104.6643
15.5594 | -104.6385
15.5633 | -104.6123
15.5671 | -104.5885
15.571 | -104.5626
15.5749 | -104.5361
15.5787 | -104.5083
15.5826 | -104.4804
15.5865 | -104.4568
15.5903 | -104.4323
15.5942 | -104.406
15.5981 | -104.3817
15.6019 | -104.3629
15.6058 | -104.3432
15.6097 | -104.3438
15.6135 | -104.3334
15.6174 | -104.3252
15.6213 | -104.3218
15.6251 | -104.3209
15.629 | -104.323
15.6328 | -104.3277
15.6367 | -104.3367
15.6406 | -104.348
15.6444 | -104.36
15.6483 | -104.3753
15.6522 | -104.3906
15.656 | -104.4094
15.6599 | -104.4293
15.6638 | -104.4491
15.6676 | -104.4713
15.6715 | -104.4935
15.6754 | -104.5125
15.6792 | -104.5311
15.6831 | -104.5488
15.687 | -104.5666
15.6908 | -104.5829
15.6947 | -104.5977
15.6986 | -104.6072
15.7024 | -104.6148
15.7063 | -104.6172
15.7101 | -104.6176
15.714 | -104.6157
15.7179 | -104.607
15.7217 | -104.5946
15.7256 | -104.5834
15.7295 | -104.5636
15.7333 | -104.5394
15.7372 | -104.5093
15.7411 | -104.4797
15.7449 | -104.4501
15.7488 | -104.4175
15.7527 | -104.3829
15.7565 | -104.3485
15.7604 | -104.3156
15.7643 | -104.2831
15.7681 | -104.251
15.772 | -104.2207
15.7758 | -104.1904
15.7797 | -104.1616
15.7836 | -104.1348
15.7874 | -104.1071
15.7913 | -104.0824
15.7952 | -104.0612
15.799 | -104.042
15.8029 | -104.0241
15.8068 | -104.0071
15.8106 | -103.9901
15.8145 | -103.978
15.8184 | -103.9694
15.8222 | -103.9585
15.8261 | -103.9425
15.83 | -103.9266
15.8338 | -103.9129
15.8377 | -103.9015
15.8415 | -103.8916
15.8454 | -103.8826
15.8493 | -103.872
15.8531 | -103.858
15.857 | -103.8437
15.8609 | -103.8319
15.8647 | -103.8208
15.8686 | -103.8082
15.8725 | -103.7976
15.8763 | -103.7879
15.8802 | -103.7763
15.8841 | -103.7641
15.8879 | -103.749
15.8918 | -103.7331
15.8957 | -103.7175
15.8995 | -103.7036
15.9034 | -103.695
15.9072 | -103.6875
15.9111 | -103.6785
15.915 | -103.6678
15.9188 | -103.6562
15.9227 | -103.6471
15.9266 | -103.6413
15.9304 | -103.6346
15.9343 | -103.6266
15.9382 | -103.6199
15.942 | -103.6135
15.9459 | -103.6081
15.9498 | -103.604
15.9536 | -103.6005
15.9575 | -103.5986
15.9614 | -103.5969
15.9652 | -103.5936
15.9691 | -103.5923
15.9729 | -103.5948
15.9768 | -103.5987
15.9807 | -103.6032
15.9845 | -103.6082
15.9884 | -103.6107
15.9923 | -103.6118
15.9961 | -103.6125
16 | -103.6125
16.0039 | -103.616
16.0077 | -103.6157
16.0116 | -103.6157
16.0155 | -103.6171
16.0193 | -103.6141
16.0232 | -103.6094
16.0271 | -103.6051
16.0309 | -103.5988
16.0348 | -103.5915
16.0386 | -103.583
16.0425 | -103.5729
16.0464 | -103.5587
16.0502 | -103.5481
16.0541 | -103.5367
16.058 | -103.5242
16.0618 | -103.5103
16.0657 | -103.4919
16.0696 | -103.4684
16.0734 | -103.4455
16.0773 | -103.4227
16.0812 | -103.4015
16.085 | -103.3788
16.0889 | -103.3571
16.0928 | -103.3356
16.0966 | -103.3165
16.1005 | -103.2986
16.1043 | -103.2809
16.1082 | -103.2663
16.1121 | -103.2517
16.1159 | -103.2411
16.1198 | -103.2317
16.1237 | -103.2232
16.1275 | -103.2148
16.1314 | -103.2043
16.1353 | -103.1963
16.1391 | -103.1915
16.143 | -103.1926
16.1469 | -103.1956
16.1507 | -103.2018
16.1546 | -103.211
16.1585 | -103.2231
16.1623 | -103.2378
16.1662 | -103.2635
16.17 | -103.2806
16.1739 | -103.3004
16.1778 | -103.3232
16.1816 | -103.3468
16.1855 | -103.3723
16.1894 | -103.398
16.1932 | -103.4233
16.1971 | -103.4513
16.201 | -103.4787
16.2048 | -103.5034
16.2087 | -103.5265
16.2126 | -103.5517
16.2164 | -103.5762
16.2203 | -103.6
16.2242 | -103.6226
16.228 | -103.6452
16.2319 | -103.6667
16.2357 | -103.6847
16.2396 | -103.7026
16.2435 | -103.7145
16.2473 | -103.7226
16.2512 | -103.732
16.2551 | -103.7387
16.2589 | -103.7437
16.2628 | -103.7475
16.2667 | -103.7505
16.2705 | -103.7507
16.2744 | -103.7466
16.2783 | -103.7444
16.2821 | -103.7391
16.286 | -103.7393
16.2899 | -103.7373
16.2937 | -103.7321
16.2976 | -103.732
16.3014 | -103.7364
16.3053 | -103.7389
16.3092 | -103.7428
16.313 | -103.748
16.3169 | -103.7547
16.3208 | -103.7599
16.3246 | -103.7647
16.3285 | -103.7723
16.3324 | -103.7806
16.3362 | -103.7912
16.3401 | -103.8054
16.344 | -103.8213
16.3478 | -103.8377
16.3517 | -103.8579
16.3556 | -103.8808
16.3594 | -103.9041
16.3633 | -103.9295
16.3671 | -103.9576
16.371 | -103.9862
16.3749 | -104.0125
16.3787 | -104.0426
16.3826 | -104.0721
16.3865 | -104.1013
16.3903 | -104.1329
16.3942 | -104.1656
16.3981 | -104.1982
16.4019 | -104.2304
16.4058 | -104.2637
16.4097 | -104.2944
16.4135 | -104.3233
16.4174 | -104.3519
16.4213 | -104.3792
16.4251 | -104.4062
16.429 | -104.434
16.4328 | -104.4598
16.4367 | -104.481
16.4406 | -104.5004
16.4444 | -104.5198
16.4483 | -104.5343
16.4522 | -104.5484
16.456 | -104.5627
16.4599 | -104.5736
16.4638 | -104.581
16.4676 | -104.5876
16.4715 | -104.5933
16.4754 | -104.5984
16.4792 | -104.6041
16.4831 | -104.6102
16.487 | -104.6151
16.4908 | -104.62
16.4947 | -104.6218
16.4986 | -104.6249
16.5024 | -104.6266
16.5063 | -104.6286
16.5101 | -104.6338
16.514 | -104.6391
16.5179 | -104.6439
16.5217 | -104.6433
16.5256 | -104.6441
16.5295 | -104.6478
16.5333 | -104.6491
16.5372 | -104.6511
16.5411 | -104.655
16.5449 | -104.6566
16.5488 | -104.6578
16.5527 | -104.6583
16.5565 | -104.657
16.5604 | -104.6591
16.5643 | -104.6589
16.5681 | -104.6584
16.572 | -104.6577
16.5758 | -104.6585
16.5797 | -104.6592
16.5836 | -104.6579
16.5874 | -104.655
16.5913 | -104.6504
16.5952 | -104.6474
16.599 | -104.6447
16.6029 | -104.6401
16.6068 | -104.6384
16.6106 | -104.6363
16.6145 | -104.6314
16.6184 | -104.6278
16.6222 | -104.6252
16.6261 | -104.623
16.63 | -104.6214
16.6338 | -104.6148
16.6377 | -104.6087
16.6415 | -104.6031
16.6454 | -104.5977
16.6493 | -104.593
16.6531 | -104.5876
16.657 | -104.5829
16.6609 | -104.5786
16.6647 | -104.5743
16.6686 | -104.5691
16.6725 | -104.5652
16.6763 | -104.564
16.6802 | -104.5642
16.6841 | -104.56
16.6879 | -104.555
16.6918 | -104.5549
16.6957 | -104.557
16.6995 | -104.5573
16.7034 | -104.5574
16.7072 | -104.5603
16.7111 | -104.5632
16.715 | -104.5662
16.7188 | -104.5671
16.7227 | -104.567
16.7266 | -104.5607
16.7304 | -104.5596
16.7343 | -104.5583
16.7382 | -104.5555
16.742 | -104.551
16.7459 | -104.5447
16.7498 | -104.5374
16.7536 | -104.5284
16.7575 | -104.5182
16.7614 | -104.5085
16.7652 | -104.4984
16.7691 | -104.4856
16.7729 | -104.4733
16.7768 | -104.462
16.7807 | -104.4486
16.7845 | -104.4338
16.7884 | -104.4205
16.7923 | -104.4061
16.7961 | -104.3897
16.8 | -104.372
16.8039 | -104.3538
16.8077 | -104.3347
16.8116 | -104.3183
16.8155 | -104.3018
16.8193 | -104.2833
16.8232 | -104.2665
16.8271 | -104.2501
16.8309 | -104.2338
16.8348 | -104.218
16.8386 | -104.1995
16.8425 | -104.1793
16.8464 | -104.1612
16.8502 | -104.1441
16.8541 | -104.1285
16.858 | -104.1071
16.8618 | -104.0852
16.8657 | -104.0638
16.8696 | -104.0441
16.8734 | -104.0285
16.8773 | -104.0071
16.8812 | -103.9857
16.885 | -103.9669
16.8889 | -103.9472
16.8928 | -103.9277
16.8966 | -103.9076
16.9005 | -103.886
16.9043 | -103.865
16.9082 | -103.8438
16.9121 | -103.8192
16.9159 | -103.7936
16.9198 | -103.768
16.9237 | -103.744
16.9275 | -103.7152
16.9314 | -103.6839
16.9353 | -103.6511
16.9391 | -103.6198
16.943 | -103.5864
16.9469 | -103.5526
16.9507 | -103.5218
16.9546 | -103.4901
16.9585 | -103.4585
16.9623 | -103.4278
16.9662 | -103.3968
16.97 | -103.3644
16.9739 | -103.3343
16.9778 | -103.3066
16.9816 | -103.2793
16.9855 | -103.2525
16.9894 | -103.2289
16.9932 | -103.2068
16.9971 | -103.1867
17.001 | -103.1685
17.0048 | -103.1488
17.0087 | -103.131
17.0126 | -103.1171
17.0164 | -103.1029
17.0203 | -103.0875
17.0242 | -103.0751
17.028 | -103.0661
17.0319 | -103.0568
17.0357 | -103.0512
17.0396 | -103.0469
17.0435 | -103.0424
17.0473 | -103.0382
17.0512 | -103.0337
17.0551 | -103.0272
17.0589 | -103.0205
17.0628 | -103.0171
17.0667 | -103.0135
17.0705 | -103.0094
17.0744 | -103.0054
17.0783 | -103.0016
17.0821 | -102.9982
17.086 | -102.9957
17.0899 | -102.9917
17.0937 | -102.9891
17.0976 | -102.987
17.1014 | -102.9888
17.1053 | -102.9866
17.1092 | -102.9843
17.113 | -102.9809
17.1169 | -102.977
17.1208 | -102.977
17.1246 | -102.972
17.1285 | -102.9649
17.1324 | -102.9604
17.1362 | -102.9534
17.1401 | -102.944
17.144 | -102.9369
17.1478 | -102.9295
17.1517 | -102.9215
17.1556 | -102.9115
17.1594 | -102.8997
17.1633 | -102.887
17.1671 | -102.875
17.171 | -102.8633
17.1749 | -102.8495
17.1787 | -102.8357
17.1826 | -102.8164
17.1865 | -102.7944
17.1903 | -102.7729
17.1942 | -102.7468
17.1981 | -102.7195
17.2019 | -102.6914
17.2058 | -102.6631
17.2097 | -102.6359
17.2135 | -102.6059
17.2174 | -102.5725
17.2213 | -102.5394
17.2251 | -102.509
17.229 | -102.4762
17.2328 | -102.4423
17.2367 | -102.4072
17.2406 | -102.3738
17.2444 | -102.3418
17.2483 | -102.3125
17.2522 | -102.2835
17.256 | -102.2548
17.2599 | -102.2281
17.2638 | -102.2058
17.2676 | -102.1847
17.2715 | -102.1637
17.2754 | -102.1449
17.2792 | -102.1271
17.2831 | -102.1207
17.287 | -102.1051
17.2908 | -102.0939
17.2947 | -102.0837
17.2986 | -102.0734
17.3024 | -102.065
17.3063 | -102.0592
17.3101 | -102.0525
17.314 | -102.0461
17.3179 | -102.042
17.3217 | -102.0394
17.3256 | -102.0356
17.3295 | -102.0315
17.3333 | -102.0261
17.3372 | -102.0191
17.3411 | -102.0119
17.3449 | -102.0077
17.3488 | -102.0022
17.3527 | -101.9957
17.3565 | -101.988
17.3604 | -101.9808
17.3643 | -101.9712
17.3681 | -101.9561
17.372 | -101.9428
17.3758 | -101.9306
17.3797 | -101.9141
17.3836 | -101.8968
17.3874 | -101.8793
17.3913 | -101.8583
17.3952 | -101.8369
17.399 | -101.8167
17.4029 | -101.7945
17.4068 | -101.7715
17.4106 | -101.7469
17.4145 | -101.7209
17.4184 | -101.697
17.4222 | -101.6716
17.4261 | -101.6449
17.43 | -101.6203
17.4338 | -101.5969
17.4377 | -101.5708
17.4415 | -101.5494
17.4454 | -101.5264
17.4493 | -101.5054
17.4531 | -101.4864
17.457 | -101.4679
17.4609 | -101.4501
17.4647 | -101.4319
17.4686 | -101.4155
17.4725 | -101.4022
17.4763 | -101.3919
17.4802 | -101.3847
17.4841 | -101.379
17.4879 | -101.3726
17.4918 | -101.3659
17.4957 | -101.3606
17.4995 | -101.3591
17.5034 | -101.3608
17.5072 | -101.3633
17.5111 | -101.3645
17.515 | -101.3664
17.5188 | -101.3701
17.5227 | -101.373
17.5266 | -101.3762
17.5304 | -101.3811
17.5343 | -101.3869
17.5382 | -101.3927
17.542 | -101.3958
17.5459 | -101.401
17.5498 | -101.4065
17.5536 | -101.412
17.5575 | -101.4165
17.5614 | -101.4206
17.5652 | -101.4241
17.5691 | -101.4261
17.5729 | -101.4245
17.5768 | -101.4255
17.5807 | -101.4279
17.5845 | -101.4324
17.5884 | -101.4349
17.5923 | -101.4361
17.5961 | -101.436
17.6 | -101.438
17.6039 | -101.4431
17.6077 | -101.4477
17.6116 | -101.451
17.6155 | -101.4565
17.6193 | -101.4637
17.6232 | -101.4709
17.6271 | -101.4788
17.6309 | -101.4889
17.6348 | -101.5006
17.6386 | -101.5129
17.6425 | -101.5269
17.6464 | -101.5396
17.6502 | -101.5512
17.6541 | -101.5657
17.658 | -101.5783
17.6618 | -101.5921
17.6657 | -101.6053
17.6696 | -101.6169
17.6734 | -101.6305
17.6773 | -101.646
17.6812 | -101.6609
17.685 | -101.6737
17.6889 | -101.6854
17.6928 | -101.6951
17.6966 | -101.7102
17.7005 | -101.7208
17.7043 | -101.7297
17.7082 | -101.7396
17.7121 | -101.7512
17.7159 | -101.7627
17.7198 | -101.7698
17.7237 | -101.7733
17.7275 | -101.7805
17.7314 | -101.7859
17.7353 | -101.7896
17.7391 | -101.7941
17.743 | -101.7986
17.7469 | -101.8026
17.7507 | -101.8034
17.7546 | -101.802
17.7585 | -101.799
17.7623 | -101.7965
17.7662 | -101.7946
17.77 | -101.7948
17.7739 | -101.7943
17.7778 | -101.7929
17.7816 | -101.7894
17.7855 | -101.7857
17.7894 | -101.7815
17.7932 | -101.7776
17.7971 | -101.7724
17.801 | -101.7696
17.8048 | -101.7648
17.8087 | -101.7583
17.8126 | -101.7533
17.8164 | -101.7491
17.8203 | -101.7446
17.8242 | -101.7424
17.828 | -101.7372
17.8319 | -101.731
17.8357 | -101.7234
17.8396 | -101.7085
17.8435 | -101.7
17.8473 | -101.6933
17.8512 | -101.6852
17.8551 | -101.6759
17.8589 | -101.6636
17.8628 | -101.6486
17.8667 | -101.6358
17.8705 | -101.6202
17.8744 | -101.6016
17.8783 | -101.584
17.8821 | -101.5652
17.886 | -101.5463
17.8899 | -101.527
17.8937 | -101.5067
17.8976 | -101.4829
17.9014 | -101.4592
17.9053 | -101.4361
17.9092 | -101.4126
17.913 | -101.3892
17.9169 | -101.3612
17.9208 | -101.33
17.9246 | -101.3018
17.9285 | -101.2714
17.9324 | -101.2425
17.9362 | -101.2139
17.9401 | -101.1858
17.944 | -101.1564
17.9478 | -101.1255
17.9517 | -101.0946
17.9556 | -101.0664
17.9594 | -101.0396
17.9633 | -101.012
17.9672 | -100.9828
17.971 | -100.9575
17.9749 | -100.9339
17.9787 | -100.9109
17.9826 | -100.888
17.9865 | -100.865
17.9903 | -100.8454
17.9942 | -100.8256
17.9981 | -100.8064
18.0019 | -100.7891
18.0058 | -100.7732
18.0097 | -100.7573
18.0135 | -100.746
18.0174 | -100.7318
18.0213 | -100.7184
18.0251 | -100.7058
18.029 | -100.6967
18.0328 | -100.6879
18.0367 | -100.6792
18.0406 | -100.6721
18.0444 | -100.6659
18.0483 | -100.6608
18.0522 | -100.6572
18.056 | -100.6527
18.0599 | -100.6516
18.0638 | -100.6529
18.0676 | -100.6548
18.0715 | -100.656
18.0754 | -100.6591
18.0792 | -100.6629
18.0831 | -100.6663
18.087 | -100.6712
18.0908 | -100.6777
18.0947 | -100.6842
18.0986 | -100.6895
18.1024 | -100.6934
18.1063 | -100.6971
18.1101 | -100.6992
18.114 | -100.7018
18.1179 | -100.7075
18.1217 | -100.715
18.1256 | -100.7182
18.1295 | -100.7239
18.1333 | -100.7324
18.1372 | -100.7344
18.1411 | -100.7336
18.1449 | -100.735
18.1488 | -100.7362
18.1527 | -100.7376
18.1565 | -100.7393
18.1604 | -100.7388
18.1643 | -100.7382
18.1681 | -100.7381
18.172 | -100.7361
18.1758 | -100.7354
18.1797 | -100.7358
18.1836 | -100.7308
18.1874 | -100.7258
18.1913 | -100.717
18.1952 | -100.7114
18.199 | -100.7053
18.2029 | -100.6995
18.2068 | -100.6928
18.2106 | -100.6839
18.2145 | -100.6761
18.2184 | -100.6639
18.2222 | -100.6501
18.2261 | -100.6383
18.23 | -100.6242
18.2338 | -100.6072
18.2377 | -100.5871
18.2415 | -100.565
18.2454 | -100.5433
18.2493 | -100.5227
18.2531 | -100.5031
18.257 | -100.4858
18.2609 | -100.4647
18.2647 | -100.4411
18.2686 | -100.4175
18.2725 | -100.395
18.2763 | -100.3724
18.2802 | -100.3443
18.2841 | -100.3174
18.2879 | -100.2895
18.2918 | -100.2607
18.2957 | -100.2342
18.2995 | -100.2056
18.3034 | -100.1766
18.3072 | -100.1481
18.3111 | -100.1194
18.315 | -100.0902
18.3188 | -100.0578
18.3227 | -100.0222
18.3266 | -99.9904
18.3304 | -99.9629
18.3343 | -99.9347
18.3382 | -99.9073
18.342 | -99.8818
18.3459 | -99.856
18.3498 | -99.8302
18.3536 | -99.807
18.3575 | -99.7893
18.3614 | -99.7726
18.3652 | -99.7541
18.3691 | -99.7398
18.3729 | -99.728
18.3768 | -99.7154
18.3807 | -99.7031
18.3845 | -99.6954
18.3884 | -99.6871
18.3923 | -99.6809
18.3961 | -99.6877
18.4 | -99.6873
18.4039 | -99.6874
18.4077 | -99.6819
18.4116 | -99.6813
18.4155 | -99.6829
18.4193 | -99.686
18.4232 | -99.6917
18.4271 | -99.6976
18.4309 | -99.7023
18.4348 | -99.708
18.4386 | -99.7134
18.4425 | -99.7175
18.4464 | -99.7223
18.4502 | -99.7271
18.4541 | -99.7338
18.458 | -99.7378
18.4618 | -99.7414
18.4657 | -99.7474
18.4696 | -99.7541
18.4734 | -99.7594
18.4773 | -99.7638
18.4812 | -99.7705
18.485 | -99.7783
18.4889 | -99.787
18.4928 | -99.7959
18.4966 | -99.8052
18.5005 | -99.8156
18.5043 | -99.8254
18.5082 | -99.8347
18.5121 | -99.8407
18.5159 | -99.848
18.5198 | -99.8554
18.5237 | -99.8651
18.5275 | -99.8752
18.5314 | -99.8842
18.5353 | -99.894
18.5391 | -99.9069
18.543 | -99.9238
18.5469 | -99.937
18.5507 | -99.9505
18.5546 | -99.9662
18.5585 | -99.9807
18.5623 | -99.9962
18.5662 | -100.0109
18.57 | -100.0244
18.5739 | -100.0372
18.5778 | -100.0519
18.5816 | -100.0684
18.5855 | -100.0869
18.5894 | -100.1005
18.5932 | -100.1121
18.5971 | -100.1246
18.601 | -100.1349
18.6048 | -100.1447
18.6087 | -100.1572
18.6126 | -100.169
18.6164 | -100.1801
18.6203 | -100.1889
18.6242 | -100.1983
18.628 | -100.2103
18.6319 | -100.2194
18.6357 | -100.2243
18.6396 | -100.2318
18.6435 | -100.2424
18.6473 | -100.2502
18.6512 | -100.2576
18.6551 | -100.2655
18.6589 | -100.2732
18.6628 | -100.2799
18.6667 | -100.2853
18.6705 | -100.2935
18.6744 | -100.306
18.6783 | -100.3124
18.6821 | -100.3168
18.686 | -100.323
18.6899 | -100.3282
18.6937 | -100.3336
18.6976 | -100.3397
18.7014 | -100.3458
18.7053 | -100.3503
18.7092 | -100.356
18.713 | -100.361
18.7169 | -100.3644
18.7208 | -100.3662
18.7246 | -100.3703
18.7285 | -100.3747
18.7324 | -100.375
18.7362 | -100.3757
18.7401 | -100.3812
18.744 | -100.385
18.7478 | -100.3889
18.7517 | -100.3953
18.7556 | -100.3967
18.7594 | -100.3996
18.7633 | -100.4062
18.7672 | -100.4141
18.771 | -100.4226
18.7749 | -100.4324
18.7787 | -100.4432
18.7826 | -100.4535
18.7865 | -100.4649
18.7903 | -100.4757
18.7942 | -100.4882
18.7981 | -100.5032
18.8019 | -100.514
18.8058 | -100.5237
18.8097 | -100.5342
18.8135 | -100.5444
18.8174 | -100.555
18.8213 | -100.5658
18.8251 | -100.5781
18.829 | -100.5905
18.8329 | -100.6026
18.8367 | -100.616
18.8406 | -100.626
18.8444 | -100.6343
18.8483 | -100.6436
18.8522 | -100.6526
18.856 | -100.6605
18.8599 | -100.663
18.8638 | -100.6647
18.8676 | -100.665
18.8715 | -100.6611
18.8754 | -100.6576
18.8792 | -100.6555
18.8831 | -100.65
18.887 | -100.6403
18.8908 | -100.6309
18.8947 | -100.6228
18.8986 | -100.6143
18.9024 | -100.6068
18.9063 | -100.5943
18.9101 | -100.5756
18.914 | -100.5576
18.9179 | -100.5377
18.9217 | -100.5156
18.9256 | -100.4959
18.9295 | -100.4772
18.9333 | -100.4553
18.9372 | -100.4312
18.9411 | -100.4083
18.9449 | -100.3847
18.9488 | -100.364
18.9527 | -100.3396
18.9565 | -100.316
18.9604 | -100.2927
18.9643 | -100.2699
18.9681 | -100.2494
18.972 | -100.2317
18.9758 | -100.2149
18.9797 | -100.1985
18.9836 | -100.1803
18.9874 | -100.1619
18.9913 | -100.1412
18.9952 | -100.1223
18.999 | -100.1063
19.0029 | -100.0931
19.0068 | -100.0805
19.0106 | -100.0653
19.0145 | -100.0527
19.0184 | -100.0416
19.0222 | -100.0311
19.0261 | -100.0204
19.03 | -100.0085
19.0338 | -99.9968
19.0377 | -99.9872
19.0415 | -99.9788
19.0454 | -99.9701
19.0493 | -99.9611
19.0531 | -99.9503
19.057 | -99.9384
19.0609 | -99.9293
19.0647 | -99.9201
19.0686 | -99.9104
19.0725 | -99.9019
19.0763 | -99.8957
19.0802 | -99.8873
19.0841 | -99.8799
19.0879 | -99.8717
19.0918 | -99.8599
19.0957 | -99.8474
19.0995 | -99.8365
19.1034 | -99.8248
19.1072 | -99.8097
19.1111 | -99.7986
19.115 | -99.7888
19.1188 | -99.7764
19.1227 | -99.7631
19.1266 | -99.7514
19.1304 | -99.7409
19.1343 | -99.7296
19.1382 | -99.7189
19.142 | -99.7096
19.1459 | -99.6965
19.1498 | -99.6827
19.1536 | -99.6723
19.1575 | -99.6608
19.1614 | -99.645
19.1652 | -99.6292
19.1691 | -99.6185
19.1729 | -99.6109
19.1768 | -99.6051
19.1807 | -99.5981
19.1845 | -99.5883
19.1884 | -99.5813
19.1923 | -99.5741
19.1961 | -99.5648
19.2 | -99.5557
19.2039 | -99.5489
19.2077 | -99.5431
19.2116 | -99.538
19.2155 | -99.5352
19.2193 | -99.5287
19.2232 | -99.5233
19.2271 | -99.5178
19.2309 | -99.5139
19.2348 | -99.5094
19.2386 | -99.5043
19.2425 | -99.4994
19.2464 | -99.4936
19.2502 | -99.4906
19.2541 | -99.4869
19.258 | -99.4856
19.2618 | -99.4853
19.2657 | -99.4869
19.2696 | -99.4874
19.2734 | -99.4895
19.2773 | -99.492
19.2812 | -99.4927
19.285 | -99.4972
19.2889 | -99.5037
19.2928 | -99.5089
19.2966 | -99.514
19.3005 | -99.5226
19.3043 | -99.5318
19.3082 | -99.5419
19.3121 | -99.5509
19.3159 | -99.5589
19.3198 | -99.5671
19.3237 | -99.5777
19.3275 | -99.5911
19.3314 | -99.6031
19.3353 | -99.6131
19.3391 | -99.6245
19.343 | -99.6352
19.3469 | -99.6452
19.3507 | -99.6534
19.3546 | -99.6621
19.3585 | -99.67
19.3623 | -99.6785
19.3662 | -99.6895
19.37 | -99.698
19.3739 | -99.7043
19.3778 | -99.7117
19.3816 | -99.7206
19.3855 | -99.7233
19.3894 | -99.7258
19.3932 | -99.7305
19.3971 | -99.7323
19.401 | -99.7319
19.4048 | -99.7336
19.4087 | -99.7333
19.4126 | -99.7331
19.4164 | -99.7321
19.4203 | -99.7307
19.4242 | -99.7295
19.428 | -99.7271
19.4319 | -99.7251
19.4357 | -99.7246
19.4396 | -99.7176
19.4435 | -99.7067
19.4473 | -99.6938
19.4512 | -99.6833
19.4551 | -99.6765
19.4589 | -99.6721
19.4628 | -99.6634
19.4667 | -99.6562
19.4705 | -99.6496
19.4744 | -99.6429
19.4783 | -99.6341
19.4821 | -99.6245
19.486 | -99.6145
19.4899 | -99.604
19.4937 | -99.5971
19.4976 | -99.593
19.5014 | -99.5902
19.5053 | -99.5872
19.5092 | -99.5828
19.513 | -99.5833
19.5169 | -99.5805
19.5208 | -99.5782
19.5246 | -99.5767
19.5285 | -99.5758
19.5324 | -99.576
19.5362 | -99.5775
19.5401 | -99.5781
19.544 | -99.5785
19.5478 | -99.5794
19.5517 | -99.5799
19.5556 | -99.5825
19.5594 | -99.5862
19.5633 | -99.5863
19.5672 | -99.5867
19.571 | -99.5895
19.5749 | -99.5938
19.5787 | -99.5995
19.5826 | -99.606
19.5865 | -99.6097
19.5903 | -99.614
19.5942 | -99.6189
19.5981 | -99.6239
19.6019 | -99.628
19.6058 | -99.6293
19.6097 | -99.6338
19.6135 | -99.6415
19.6174 | -99.6466
19.6213 | -99.6508
19.6251 | -99.6571
19.629 | -99.6628
19.6329 | -99.6683
19.6367 | -99.674
19.6406 | -99.6788
19.6444 | -99.6849
19.6483 | -99.6931
19.6522 | -99.7011
19.656 | -99.7087
19.6599 | -99.7168
19.6638 | -99.7284
19.6676 | -99.7376
19.6715 | -99.7454
19.6754 | -99.7555
19.6792 | -99.765
19.6831 | -99.7735
19.687 | -99.7832
19.6908 | -99.793
19.6947 | -99.8004
19.6986 | -99.8065
19.7024 | -99.8104
19.7063 | -99.8156
19.7101 | -99.8212
19.714 | -99.8262
19.7179 | -99.8342
19.7217 | -99.8414
19.7256 | -99.848
19.7295 | -99.8521
19.7333 | -99.8557
19.7372 | -99.8599
19.7411 | -99.8676
19.7449 | -99.8732
19.7488 | -99.878
19.7527 | -99.883
19.7565 | -99.8898
19.7604 | -99.8984
19.7643 | -99.9063
19.7681 | -99.9156
19.772 | -99.9246
19.7758 | -99.9308
19.7797 | -99.939
19.7836 | -99.9474
19.7874 | -99.9574
19.7913 | -99.9673
19.7952 | -99.9782
19.799 | -99.9888
19.8029 | -100.0005
19.8068 | -100.013
19.8106 | -100.0233
19.8145 | -100.0337
19.8184 | -100.0431
19.8222 | -100.0542
19.8261 | -100.0671
19.83 | -100.0798
19.8338 | -100.0901
19.8377 | -100.0977
19.8415 | -100.1069
19.8454 | -100.1161
19.8493 | -100.1252
19.8531 | -100.1346
19.857 | -100.1419
19.8609 | -100.1498
19.8647 | -100.1559
19.8686 | -100.1637
19.8725 | -100.1693
19.8763 | -100.1756
19.8802 | -100.1833
19.8841 | -100.1895
19.8879 | -100.1984
19.8918 | -100.2043
19.8957 | -100.2076
19.8995 | -100.2093
19.9034 | -100.2115
19.9072 | -100.212
19.9111 | -100.213
19.915 | -100.212
19.9188 | -100.2105
19.9227 | -100.2086
19.9266 | -100.2058
19.9304 | -100.2001
19.9343 | -100.1945
19.9382 | -100.1877
19.942 | -100.1807
19.9459 | -100.1728
19.9498 | -100.1629
19.9536 | -100.1553
19.9575 | -100.1505
19.9614 | -100.146
19.9652 | -100.1393
19.9691 | -100.1334
19.9729 | -100.1284
19.9768 | -100.1259
19.9807 | -100.1252
19.9845 | -100.1268
19.9884 | -100.1276
19.9923 | -100.1286
19.9961 | -100.13
20 | -100.133
20.0039 | -100.1367
20.0077 | -100.1419
20.0116 | -100.1466
20.0155 | -100.1516
20.0193 | -100.1566
20.0232 | -100.1617
20.0271 | -100.1682
20.0309 | -100.1778
20.0348 | -100.1909
20.0386 | -100.2045
20.0425 | -100.2167
20.0464 | -100.2298
20.0502 | -100.2412
20.0541 | -100.2533
20.058 | -100.2657
20.0618 | -100.2783
20.0657 | -100.2918
20.0696 | -100.3025
20.0734 | -100.315
20.0773 | -100.325
20.0812 | -100.3372
20.085 | -100.3506
20.0889 | -100.3619
20.0928 | -100.3722
20.0966 | -100.3832
20.1005 | -100.3946
20.1043 | -100.4052
20.1082 | -100.416
20.1121 | -100.4244
20.1159 | -100.4296
20.1198 | -100.4368
20.1237 | -100.4435
20.1275 | -100.4489
20.1314 | -100.4543
20.1353 | -100.4586
20.1391 | -100.4621
20.143 | -100.4656
20.1469 | -100.4674
20.1507 | -100.4703
20.1546 | -100.4728
20.1585 | -100.4761
20.1623 | -100.479
20.1662 | -100.4793
20.17 | -100.4812
20.1739 | -100.4802
20.1778 | -100.478
20.1816 | -100.4785
20.1855 | -100.4796
20.1894 | -100.4795
20.1932 | -100.4777
20.1971 | -100.4784
20.201 | -100.4783
20.2048 | -100.4771
20.2087 | -100.4768
20.2126 | -100.4747
20.2164 | -100.474
20.2203 | -100.4731
20.2242 | -100.4711
20.228 | -100.4694
20.2319 | -100.4678
20.2357 | -100.4658
20.2396 | -100.4553
20.2435 | -100.4451
20.2473 | -100.4399
20.2512 | -100.4345
20.2551 | -100.4271
20.2589 | -100.4187
20.2628 | -100.4084
20.2667 | -100.3971
20.2705 | -100.3851
20.2744 | -100.3744
20.2783 | -100.3634
20.2821 | -100.3521
20.286 | -100.339
20.2899 | -100.326
20.2937 | -100.313
20.2976 | -100.3011
20.3014 | -100.2893
20.3053 | -100.2772
20.3092 | -100.2643
20.313 | -100.2521
20.3169 | -100.2405
20.3208 | -100.2272
20.3246 | -100.2146
20.3285 | -100.2029
20.3324 | -100.1924
20.3362 | -100.1821
20.3401 | -100.1708
20.344 | -100.1627
20.3478 | -100.1534
20.3517 | -100.145
20.3556 | -100.1366
20.3594 | -100.1289
20.3633 | -100.1243
20.3671 | -100.1209
20.371 | -100.1154
20.3749 | -100.1102
20.3787 | -100.1064
20.3826 | -100.1049
20.3865 | -100.1044
20.3903 | -100.1049
20.3942 | -100.1031
20.3981 | -100.1008
20.4019 | -100.1029
20.4058 | -100.1054
20.4097 | -100.1072
20.4135 | -100.1099
20.4174 | -100.1104
20.4213 | -100.1108
20.4251 | -100.11
20.429 | -100.1077
20.4328 | -100.1057
20.4367 | -100.1038
20.4406 | -100.1011
20.4444 | -100.0957
20.4483 | -100.0884
20.4522 | -100.0771
20.456 | -100.0659
20.4599 | -100.0548
20.4638 | -100.0449
20.4676 | -100.032
20.4715 | -100.0158
20.4754 | -100.0022
20.4792 | -99.9855
20.4831 | -99.9646
20.487 | -99.9413
20.4908 | -99.9183
20.4947 | -99.8953
20.4986 | -99.87
20.5024 | -99.8449
20.5063 | -99.8175
20.5101 | -99.7906
20.514 | -99.7629
20.5179 | -99.7308
20.5217 | -99.6984
20.5256 | -99.6703
20.5295 | -99.6441
20.5333 | -99.6133
20.5372 | -99.58
20.5411 | -99.5465
20.5449 | -99.5146
20.5488 | -99.4842
20.5527 | -99.4537
20.5565 | -99.4236
20.5604 | -99.3929
20.5643 | -99.3626
20.5681 | -99.3353
20.572 | -99.3065
20.5758 | -99.2746
20.5797 | -99.2417
20.5836 | -99.2143
20.5874 | -99.1907
20.5913 | -99.1678
20.5952 | -99.1454
20.599 | -99.1239
20.6029 | -99.1024
20.6068 | -99.0798
20.6106 | -99.0597
20.6145 | -99.0414
20.6184 | -99.0258
20.6222 | -99.0112
20.6261 | -99.0008
20.63 | -98.9862
20.6338 | -98.9716
20.6377 | -98.958
20.6415 | -98.947
20.6454 | -98.9329
20.6493 | -98.9217
20.6531 | -98.9104
20.657 | -98.8971
20.6609 | -98.8834
20.6647 | -98.8707
20.6686 | -98.8592
20.6725 | -98.8487
20.6763 | -98.8368
20.6802 | -98.825
20.6841 | -98.8121
20.6879 | -98.8008
20.6918 | -98.7902
20.6957 | -98.7774
20.6995 | -98.7654
20.7034 | -98.7539
20.7072 | -98.7419
20.7111 | -98.7306
20.715 | -98.7193
20.7188 | -98.7093
20.7227 | -98.6977
20.7266 | -98.6837
20.7304 | -98.6714
20.7343 | -98.6598
20.7382 | -98.6504
20.742 | -98.6422
20.7459 | -98.6358
20.7498 | -98.6277
20.7536 | -98.6157
20.7575 | -98.6018
20.7614 | -98.5934
20.7652 | -98.5863
20.7691 | -98.5784
20.7729 | -98.5721
20.7768 | -98.5675
20.7807 | -98.5589
20.7845 | -98.5535
20.7884 | -98.551
20.7923 | -98.5486
20.7961 | -98.5449
20.8 | -98.5425
20.8039 | -98.54
20.8077 | -98.5377
20.8116 | -98.5394
20.8155 | -98.5406
20.8193 | -98.5398
20.8232 | -98.5409
20.8271 | -98.5425
20.8309 | -98.5445
20.8348 | -98.5488
20.8386 | -98.5521
20.8425 | -98.5558
20.8464 | -98.562
20.8502 | -98.568
20.8541 | -98.5741
20.858 | -98.579
20.8618 | -98.5837
20.8657 | -98.59
20.8696 | -98.599
20.8734 | -98.6072
20.8773 | -98.6111
20.8812 | -98.6125
20.885 | -98.616
20.8889 | -98.6214
20.8928 | -98.6279
20.8966 | -98.6363
20.9005 | -98.6466
20.9043 | -98.6539
20.9082 | -98.6599
20.9121 | -98.6645
20.9159 | -98.6695
20.9198 | -98.6763
20.9237 | -98.6818
20.9275 | -98.6868
20.9314 | -98.6925
20.9353 | -98.6965
20.9391 | -98.6979
20.943 | -98.7026
20.9469 | -98.7045
20.9507 | -98.7069
20.9546 | -98.709
20.9585 | -98.7117
20.9623 | -98.7091
20.9662 | -98.7071
20.97 | -98.7055
20.9739 | -98.7034
20.9778 | -98.6977
20.9816 | -98.6946
20.9855 | -98.6916
20.9894 | -98.6875
20.9932 | -98.6826
20.9971 | -98.6761
21.001 | -98.668
21.0048 | -98.6613
21.0087 | -98.6539
21.0126 | -98.6453
21.0164 | -98.635
21.0203 | -98.6268
21.0242 | -98.6185
21.028 | -98.608
21.0319 | -98.5969
21.0357 | -98.587
21.0396 | -98.5771
21.0435 | -98.5684
21.0473 | -98.5609
21.0512 | -98.5536
21.0551 | -98.5476
21.0589 | -98.5397
21.0628 | -98.5279
21.0667 | -98.5155
21.0705 | -98.5094
21.0744 | -98.5056
21.0783 | -98.5065
21.0821 | -98.5032
21.086 | -98.4988
21.0899 | -98.4968
21.0937 | -98.4977
21.0976 | -98.4983
21.1014 | -98.4995
21.1053 | -98.5033
21.1092 | -98.5067
21.113 | -98.5115
21.1169 | -98.5164
21.1208 | -98.5204
21.1246 | -98.5233
21.1285 | -98.5239
21.1324 | -98.5248
21.1362 | -98.5269
21.1401 | -98.5309
21.144 | -98.5339
21.1478 | -98.5357
21.1517 | -98.5393
21.1556 | -98.5422
21.1594 | -98.5453
21.1633 | -98.5483
21.1671 | -98.5511
21.171 | -98.5511
21.1749 | -98.5488
21.1787 | -98.5473
21.1826 | -98.5421
21.1865 | -98.5332
21.1903 | -98.5282
21.1942 | -98.5201
21.1981 | -98.51
21.2019 | -98.5005
21.2058 | -98.4943
21.2097 | -98.4891
21.2135 | -98.4829
21.2174 | -98.4736
21.2213 | -98.4591
21.2251 | -98.446
21.229 | -98.4357
21.2328 | -98.4264
21.2367 | -98.416
21.2406 | -98.4038
21.2444 | -98.3916
21.2483 | -98.3807
21.2522 | -98.3682
21.256 | -98.3578
21.2599 | -98.3491
21.2638 | -98.3425
21.2676 | -98.3313
21.2715 | -98.3184
21.2754 | -98.3087
21.2792 | -98.3029
21.2831 | -98.2965
21.287 | -98.2873
21.2908 | -98.2778
21.2947 | -98.2706
21.2986 | -98.2663
21.3024 | -98.2632
21.3063 | -98.2587
21.3101 | -98.2565
21.314 | -98.2549
21.3179 | -98.2538
21.3217 | -98.2538
21.3256 | -98.2556
21.3295 | -98.2558
21.3333 | -98.2551
21.3372 | -98.2529
21.3411 | -98.2516
21.3449 | -98.251
21.3488 | -98.2517
21.3527 | -98.2547
21.3565 | -98.258
21.3604 | -98.2602
21.3643 | -98.2656
21.3681 | -98.2713
21.372 | -98.2719
21.3758 | -98.271
21.3797 | -98.271
21.3836 | -98.2718
21.3874 | -98.2723
21.3913 | -98.2707
21.3952 | -98.2691
21.399 | -98.2688
21.4029 | -98.2707
21.4068 | -98.2726
21.4106 | -98.2757
21.4145 | -98.2794
21.4184 | -98.2836
21.4222 | -98.2897
21.4261 | -98.2993
21.43 | -98.3101
21.4338 | -98.319
21.4377 | -98.328
21.4415 | -98.3348
21.4454 | -98.3419
21.4493 | -98.3523
21.4531 | -98.3624
21.457 | -98.3743
21.4609 | -98.3856
21.4647 | -98.3969
21.4686 | -98.4084
21.4725 | -98.4224
21.4763 | -98.4379
21.4802 | -98.4516
21.4841 | -98.4682
21.4879 | -98.4829
21.4918 | -98.4941
21.4957 | -98.5081
21.4995 | -98.5255
21.5034 | -98.5427
21.5072 | -98.5576
21.5111 | -98.5725
21.515 | -98.5863
21.5188 | -98.5981
21.5227 | -98.6131
21.5266 | -98.6332
21.5304 | -98.6507
21.5343 | -98.6647
21.5382 | -98.6748
21.542 | -98.6856
21.5459 | -98.6997
21.5498 | -98.7125
21.5536 | -98.7246
21.5575 | -98.7368
21.5614 | -98.7474
21.5652 | -98.7556
21.5691 | -98.7663
21.5729 | -98.7794
21.5768 | -98.7911
21.5807 | -98.8002
21.5845 | -98.8071
21.5884 | -98.8123
21.5923 | -98.821
21.5961 | -98.8283
21.6 | -98.8319
21.6039 | -98.8345
21.6077 | -98.8362
21.6116 | -98.8377
21.6155 | -98.8352
21.6193 | -98.8328
21.6232 | -98.8277
21.6271 | -98.8208
21.6309 | -98.8163
21.6348 | -98.8123
21.6386 | -98.8083
21.6425 | -98.8055
21.6464 | -98.8015
21.6502 | -98.7935
21.6541 | -98.7825
21.658 | -98.7708
21.6618 | -98.7606
21.6657 | -98.7512
21.6696 | -98.7407
21.6734 | -98.7303
21.6773 | -98.7205
21.6812 | -98.7121
21.685 | -98.7038
21.6889 | -98.697
21.6928 | -98.689
21.6966 | -98.6844
21.7005 | -98.6815
21.7043 | -98.6782
21.7082 | -98.6736
21.7121 | -98.6669
21.7159 | -98.6599
21.7198 | -98.6535
21.7237 | -98.6513
21.7275 | -98.65
21.7314 | -98.6492
21.7353 | -98.6488
21.7391 | -98.6489
21.743 | -98.6544
21.7469 | -98.6577
21.7507 | -98.6601
21.7546 | -98.6645
21.7585 | -98.6681
21.7623 | -98.6709
21.7662 | -98.6749
21.77 | -98.6799
21.7739 | -98.6869
21.7778 | -98.693
21.7816 | -98.6995
21.7855 | -98.7047
21.7894 | -98.7114
21.7932 | -98.7165
21.7971 | -98.7212
21.801 | -98.7301
21.8048 | -98.7366
21.8087 | -98.7423
21.8126 | -98.7484
21.8164 | -98.7548
21.8203 | -98.7598
21.8242 | -98.7664
21.828 | -98.7753
21.8319 | -98.782
21.8357 | -98.7882
21.8396 | -98.7939
21.8435 | -98.7996
21.8473 | -98.8069
21.8512 | -98.8123
21.8551 | -98.8155
21.8589 | -98.8197
21.8628 | -98.8255
21.8667 | -98.8313
21.8705 | -98.8361
21.8744 | -98.8421
21.8783 | -98.8492
21.8821 | -98.8548
21.886 | -98.8609
21.8899 | -98.8679
21.8937 | -98.877
21.8976 | -98.8873
21.9014 | -98.896
21.9053 | -98.9054
21.9092 | -98.9175
21.913 | -98.9296
21.9169 | -98.9417
21.9208 | -98.9538
21.9246 | -98.9687
21.9285 | -98.9831
21.9324 | -98.9973
21.9362 | -99.0114
21.9401 | -99.0234
21.944 | -99.0369
21.9478 | -99.0512
21.9517 | -99.0643
21.9556 | -99.0772
21.9594 | -99.0915
21.9633 | -99.1073
21.9672 | -99.1242
21.971 | -99.1385
21.9749 | -99.1521
21.9787 | -99.1674
21.9826 | -99.1821
21.9865 | -99.1963
21.9903 | -99.2096
21.9942 | -99.2217
21.9981 | -99.2345
22.0019 | -99.2488
22.0058 | -99.2621
22.0097 | -99.2741
22.0135 | -99.287
22.0174 | -99.2993
22.0213 | -99.3139
22.0251 | -99.3272
22.029 | -99.3403
22.0328 | -99.3505
22.0367 | -99.36
22.0406 | -99.3699
22.0444 | -99.3806
22.0483 | -99.39
22.0522 | -99.3999
22.056 | -99.4075
22.0599 | -99.4157
22.0638 | -99.4261
22.0676 | -99.4352
22.0715 | -99.4437
22.0754 | -99.454
22.0792 | -99.4639
22.0831 | -99.4723
22.087 | -99.4776
22.0908 | -99.4871
22.0947 | -99.4911
22.0986 | -99.4997
22.1024 | -99.5085
22.1063 | -99.5172
22.1101 | -99.5254
22.114 | -99.5319
22.1179 | -99.5368
22.1217 | -99.5393
22.1256 | -99.5433
22.1295 | -99.5472
22.1333 | -99.5514
22.1372 | -99.5558
22.1411 | -99.5581
22.1449 | -99.5614
22.1488 | -99.5636
22.1527 | -99.5669
22.1565 | -99.5689
22.1604 | -99.5693
22.1643 | -99.5721
22.1681 | -99.5755
22.172 | -99.5766
22.1758 | -99.5759
22.1797 | -99.5775
22.1836 | -99.5799
22.1874 | -99.5822
22.1913 | -99.5836
22.1952 | -99.5876
22.199 | -99.5924
22.2029 | -99.5986
22.2068 | -99.6055
22.2106 | -99.6121
22.2145 | -99.6207
22.2184 | -99.6318
22.2222 | -99.6431
22.2261 | -99.6564
22.23 | -99.6719
22.2338 | -99.6891
22.2377 | -99.7078
22.2415 | -99.7278
22.2454 | -99.7498
22.2493 | -99.7725
22.2531 | -99.7955
22.257 | -99.8196
22.2609 | -99.845
22.2647 | -99.8715
22.2686 | -99.9002
22.2725 | -99.9318
22.2763 | -99.9648
22.2802 | -99.9973
22.2841 | -100.0307
22.2879 | -100.0641
22.2918 | -100.0972
22.2957 | -100.132
22.2995 | -100.1647
22.3034 | -100.1978
22.3072 | -100.2299
22.3111 | -100.2623
22.315 | -100.2938
22.3188 | -100.3244
22.3227 | -100.3537
22.3266 | -100.3826
22.3304 | -100.4105
22.3343 | -100.4369
22.3382 | -100.4636
22.342 | -100.49
22.3459 | -100.5146
22.3498 | -100.5375
22.3536 | -100.5598
22.3575 | -100.5808
22.3614 | -100.5992
22.3652 | -100.6193
22.3691 | -100.6381
22.3729 | -100.6554
22.3768 | -100.6669
22.3807 | -100.6775
22.3845 | -100.6882
22.3884 | -100.6992
22.3923 | -100.7071
22.3961 | -100.7121
22.4 | -100.7156
22.4039 | -100.7222
22.4077 | -100.7257
22.4116 | -100.7336
22.4155 | -100.733
22.4193 | -100.7324
22.4232 | -100.7306
22.4271 | -100.7291
22.4309 | -100.7272
22.4348 | -100.7244
22.4386 | -100.7236
22.4425 | -100.7215
22.4464 | -100.7173
22.4502 | -100.7129
22.4541 | -100.7036
22.458 | -100.6961
22.4618 | -100.6863
22.4657 | -100.6744
22.4696 | -100.6599
22.4734 | -100.6494
22.4773 | -100.6399
22.4812 | -100.6277
22.485 | -100.6122
22.4889 | -100.5958
22.4928 | -100.5764
22.4966 | -100.5549
22.5005 | -100.5355
22.5043 | -100.5151
22.5082 | -100.4939
22.5121 | -100.4731
22.5159 | -100.4517
22.5198 | -100.4284
22.5237 | -100.4043
22.5275 | -100.3794
22.5314 | -100.3536
22.5353 | -100.3269
22.5391 | -100.2989
22.543 | -100.2698
22.5469 | -100.2403
22.5507 | -100.2107
22.5546 | -100.1802
22.5585 | -100.1491
22.5623 | -100.1197
22.5662 | -100.0912
22.57 | -100.0622
22.5739 | -100.033
22.5778 | -100.0049
22.5816 | -99.9783
22.5855 | -99.9531
22.5894 | -99.9285
22.5932 | -99.9062
22.5971 | -99.8862
22.601 | -99.8678
22.6048 | -99.8516
22.6087 | -99.8381
22.6126 | -99.8256
22.6164 | -99.8155
22.6203 | -99.8093
22.6242 | -99.8065
22.628 | -99.8073
22.6319 | -99.8101
22.6357 | -99.8149
22.6396 | -99.8215
22.6435 | -99.8311
22.6473 | -99.843
22.6512 | -99.8585
22.6551 | -99.8772
22.6589 | -99.8971
22.6628 | -99.92
22.6667 | -99.9443
22.6705 | -99.9709
22.6744 | -99.9989
22.6783 | -100.0286
22.6821 | -100.0594
22.686 | -100.0913
22.6899 | -100.1229
22.6937 | -100.155
22.6976 | -100.1881
22.7014 | -100.2205
22.7053 | -100.2514
22.7092 | -100.2788
22.713 | -100.3079
22.7169 | -100.3382
22.7208 | -100.3678
22.7246 | -100.3915
22.7285 | -100.4132
22.7324 | -100.4336
22.7362 | -100.4538
22.7401 | -100.4695
22.744 | -100.4864
22.7478 | -100.5016
22.7517 | -100.5128
22.7556 | -100.5246
22.7594 | -100.5331
22.7633 | -100.5367
22.7672 | -100.5383
22.771 | -100.5383
22.7749 | -100.5381
22.7787 | -100.5355
22.7826 | -100.5319
22.7865 | -100.5263
22.7903 | -100.5198
22.7942 | -100.5072
22.7981 | -100.4926
22.8019 | -100.4771
22.8058 | -100.4603
22.8097 | -100.4425
22.8135 | -100.4252
22.8174 | -100.4078
22.8213 | -100.3857
22.8251 | -100.364
22.829 | -100.3417
22.8329 | -100.3191
22.8367 | -100.2984
22.8406 | -100.2769
22.8444 | -100.2556
22.8483 | -100.2355
22.8522 | -100.2175
22.856 | -100.1988
22.8599 | -100.186
22.8638 | -100.1699
22.8676 | -100.1529
22.8715 | -100.1375
22.8754 | -100.1235
22.8792 | -100.1103
22.8831 | -100.0982
22.887 | -100.0874
22.8908 | -100.0787
22.8947 | -100.07
22.8986 | -100.0623
22.9024 | -100.0544
22.9063 | -100.0471
22.9101 | -100.0397
22.914 | -100.0336
22.9179 | -100.0284
22.9217 | -100.0235
22.9256 | -100.0212
22.9295 | -100.0183
22.9333 | -100.014
22.9372 | -100.0117
22.9411 | -100.0121
22.9449 | -100.0118
22.9488 | -100.0092
22.9527 | -100.0052
22.9565 | -99.9987
22.9604 | -99.9953
22.9643 | -99.9949
22.9681 | -99.9944
22.972 | -99.9939
22.9758 | -99.9904
22.9797 | -99.9879
22.9836 | -99.9864
22.9874 | -99.9829
22.9913 | -99.98
22.9952 | -99.9763
22.999 | -99.9708
23.0029 | -99.9661
23.0068 | -99.9619
23.0106 | -99.9565
23.0145 | -99.9494
23.0184 | -99.9409
23.0222 | -99.9317
23.0261 | -99.9219
23.03 | -99.9131
23.0338 | -99.9051
23.0377 | -99.895
23.0415 | -99.8796
23.0454 | -99.8642
23.0493 | -99.8514
23.0531 | -99.8385
23.057 | -99.8254
23.0609 | -99.8111
23.0647 | -99.7955
23.0686 | -99.7738
23.0725 | -99.7507
23.0763 | -99.7299
23.0802 | -99.7103
23.0841 | -99.6904
23.0879 | -99.67
23.0918 | -99.6494
23.0957 | -99.6279
23.0995 | -99.6065
23.1034 | -99.5858
23.1072 | -99.5651
23.1111 | -99.5402
23.115 | -99.5134
23.1188 | -99.4868
23.1227 | -99.458
23.1266 | -99.4292
23.1304 | -99.4009
23.1343 | -99.3718
23.1382 | -99.3449
23.142 | -99.3188
23.1459 | -99.2929
23.1498 | -99.2673
23.1536 | -99.2411
23.1575 | -99.2117
23.1614 | -99.1833
23.1652 | -99.1533
23.1691 | -99.1237
23.1729 | -99.0972
23.1768 | -99.0709
23.1807 | -99.0451
23.1845 | -99.0204
23.1884 | -98.9938
23.1923 | -98.9678
23.1961 | -98.9408
23.2 | -98.9134
23.2039 | -98.8869
23.2077 | -98.8621
23.2116 | -98.8374
23.2155 | -98.8121
23.2193 | -98.7867
23.2232 | -98.7624
23.2271 | -98.7377
23.2309 | -98.7146
23.2348 | -98.6925
23.2386 | -98.6712
23.2425 | -98.6487
23.2464 | -98.6278
23.2502 | -98.6048
23.2541 | -98.5812
23.258 | -98.5594
23.2618 | -98.5403
23.2657 | -98.5224
23.2696 | -98.5039
23.2734 | -98.4872
23.2773 | -98.4736
23.2812 | -98.4606
23.285 | -98.448
23.2889 | -98.4321
23.2928 | -98.4163
23.2966 | -98.4026
23.3005 | -98.3904
23.3043 | -98.381
23.3082 | -98.3726
23.3121 | -98.3641
23.3159 | -98.3563
23.3198 | -98.3496
23.3237 | -98.3455
23.3275 | -98.342
23.3314 | -98.335
23.3353 | -98.3294
23.3391 | -98.3265
23.343 | -98.3237
23.3469 | -98.3216
23.3507 | -98.319
23.3546 | -98.3172
23.3585 | -98.3156
23.3623 | -98.3134
23.3662 | -98.3124
23.37 | -98.3149
23.3739 | -98.3156
23.3778 | -98.3156
23.3816 | -98.316
23.3855 | -98.3164
23.3894 | -98.3169
23.3932 | -98.318
23.3971 | -98.3214
23.401 | -98.3234
23.4048 | -98.3232
23.4087 | -98.3226
23.4126 | -98.3235
23.4164 | -98.3294
23.4203 | -98.3323
23.4242 | -98.3339
23.428 | -98.3326
23.4319 | -98.3326
23.4357 | -98.3322
23.4396 | -98.3335
23.4435 | -98.3398
23.4473 | -98.3468
23.4512 | -98.3538
23.4551 | -98.3631
23.4589 | -98.3712
23.4628 | -98.3787
23.4667 | -98.385
23.4705 | -98.3897
23.4744 | -98.3949
23.4783 | -98.4018
23.4821 | -98.4074
23.486 | -98.4135
23.4899 | -98.4204
23.4937 | -98.4247
23.4976 | -98.4305
23.5014 | -98.437
23.5053 | -98.4422
23.5092 | -98.4457
23.513 | -98.4498
23.5169 | -98.4541
23.5208 | -98.4568
23.5246 | -98.4572
23.5285 | -98.4595
23.5324 | -98.4613
23.5362 | -98.462
23.5401 | -98.4609
23.544 | -98.4598
23.5478 | -98.4583
23.5517 | -98.4542
23.5556 | -98.4509
23.5594 | -98.4465
23.5633 | -98.4414
23.5672 | -98.435
23.571 | -98.4268
23.5749 | -98.4192
23.5787 | -98.4112
23.5826 | -98.4056
23.5865 | -98.3985
23.5903 | -98.3904
23.5942 | -98.3771
23.5981 | -98.3625
23.6019 | -98.3498
23.6058 | -98.3371
23.6097 | -98.3224
23.6135 | -98.3078
23.6174 | -98.2943
23.6213 | -98.2797
23.6251 | -98.2643
23.629 | -98.248
23.6329 | -98.2314
23.6367 | -98.2125
23.6406 | -98.1932
23.6444 | -98.1751
23.6483 | -98.1593
23.6522 | -98.1443
23.656 | -98.1281
23.6599 | -98.1137
23.6638 | -98.1018
23.6676 | -98.0887
23.6715 | -98.0737
23.6754 | -98.0593
23.6792 | -98.0443
23.6831 | -98.0298
23.687 | -98.0147
23.6908 | -97.9993
23.6947 | -97.986
23.6986 | -97.9733
23.7024 | -97.9588
23.7063 | -97.945
23.7101 | -97.9307
23.714 | -97.9162
23.7179 | -97.9047
23.7217 | -97.8934
23.7256 | -97.8836
23.7295 | -97.8769
23.7333 | -97.8724
23.7372 | -97.8695
23.7411 | -97.8646
23.7449 | -97.8559
23.7488 | -97.8473
23.7527 | -97.8388
23.7565 | -97.8296
23.7604 | -97.8222
23.7643 | -97.8172
23.7681 | -97.8107
23.772 | -97.805
23.7758 | -97.7987
23.7797 | -97.7901
23.7836 | -97.7806
23.7874 | -97.7688
23.7913 | -97.7578
23.7952 | -97.7473
23.799 | -97.7372
23.8029 | -97.7278
23.8068 | -97.7178
23.8106 | -97.7067
23.8145 | -97.6963
23.8184 | -97.6862
23.8222 | -97.6753
23.8261 | -97.6658
23.83 | -97.658
23.8338 | -97.6504
23.8377 | -97.6414
23.8415 | -97.6331
23.8454 | -97.6257
23.8493 | -97.6179
23.8531 | -97.61
23.857 | -97.6042
23.8609 | -97.6002
23.8647 | -97.5979
23.8686 | -97.5966
23.8725 | -97.5956
23.8763 | -97.5973
23.8802 | -97.5995
23.8841 | -97.6029
23.8879 | -97.6085
23.8918 | -97.6141
23.8957 | -97.6203
23.8995 | -97.6268
23.9034 | -97.6363
23.9072 | -97.6469
23.9111 | -97.6565
23.915 | -97.6676
23.9188 | -97.6816
23.9227 | -97.6954
23.9266 | -97.7118
23.9304 | -97.731
23.9343 | -97.7488
23.9382 | -97.7654
23.942 | -97.782
23.9459 | -97.8
23.9498 | -97.8175
23.9536 | -97.8338
23.9575 | -97.852
23.9614 | -97.8711
23.9652 | -97.8895
23.9691 | -97.9074
23.9729 | -97.9197
23.9768 | -97.9351
23.9807 | -97.9483
23.9845 | -97.9585
23.9884 | -97.9654
23.9923 | -97.9728
23.9961 | -97.9793
24 | -97.9849
24.0039 | -97.9904
24.0077 | -97.9957
24.0116 | -97.9995
24.0155 | -98.0045
24.0193 | -98.0092
24.0232 | -98.01
24.0271 | -98.0087
24.0309 | -98.0064
24.0348 | -98.0027
24.0386 | -97.9984
24.0425 | -97.9935
24.0464 | -97.9893
24.0502 | -97.9842
24.0541 | -97.976
24.058 | -97.9667
24.0618 | -97.9575
24.0657 | -97.9457
24.0696 | -97.9319
24.0734 | -97.9192
24.0773 | -97.9055
24.0812 | -97.8931
24.085 | -97.8832
24.0889 | -97.874
24.0928 | -97.8654
24.0966 | -97.8569
24.1005 | -97.8491
24.1043 | -97.8419
24.1082 | -97.8339
24.1121 | -97.8243
24.1159 | -97.8133
24.1198 | -97.8034
24.1237 | -97.7961
24.1275 | -97.7916
24.1314 | -97.7868
24.1353 | -97.7821
24.1391 | -97.7787
24.143 | -97.7767
24.1469 | -97.7753
24.1507 | -97.773
24.1546 | -97.7715
24.1585 | -97.7725
24.1623 | -97.7742
24.1662 | -97.776
24.17 | -97.7786
24.1739 | -97.7811
24.1778 | -97.7828
24.1816 | -97.783
24.1855 | -97.7853
24.1894 | -97.7884
24.1932 | -97.7923
24.1971 | -97.799
24.201 | -97.8061
24.2048 | -97.811
24.2087 | -97.8149
24.2126 | -97.8195
24.2164 | -97.8236
24.2203 | -97.8269
24.2242 | -97.83
24.228 | -97.8327
24.2319 | -97.8366
24.2357 | -97.8408
24.2396 | -97.8446
24.2435 | -97.8477
24.2473 | -97.8502
24.2512 | -97.8543
24.2551 | -97.8571
24.2589 | -97.8556
24.2628 | -97.8553
24.2667 | -97.8527
24.2705 | -97.85
24.2744 | -97.847
24.2783 | -97.8406
24.2821 | -97.8311
24.286 | -97.8197
24.2899 | -97.8071
24.2937 | -97.7942
24.2976 | -97.7807
24.3014 | -97.7656
24.3053 | -97.7496
24.3092 | -97.7335
24.313 | -97.7171
24.3169 | -97.7011
24.3208 | -97.6813
24.3246 | -97.6607
24.3285 | -97.642
24.3324 | -97.6238
24.3362 | -97.6022
24.3401 | -97.5815
24.344 | -97.5588
24.3478 | -97.5328
24.3517 | -97.5051
24.3556 | -97.4783
24.3594 | -97.4481
24.3633 | -97.4184
24.3671 | -97.3928
24.371 | -97.3677
24.3749 | -97.3416
24.3787 | -97.3152
24.3826 | -97.2869
24.3865 | -97.2587
24.3903 | -97.2296
24.3942 | -97.2009
24.3981 | -97.1752
24.4019 | -97.1498
24.4058 | -97.1245
24.4097 | -97.1012
24.4135 | -97.079
24.4174 | -97.0573
24.4213 | -97.0363
24.4251 | -97.0164
24.429 | -96.9976
24.4328 | -96.9797
24.4367 | -96.9645
24.4406 | -96.9509
24.4444 | -96.9381
24.4483 | -96.925
24.4522 | -96.9145
24.456 | -96.9077
24.4599 | -96.9039
24.4638 | -96.9016
24.4676 | -96.8989
24.4715 | -96.8957
24.4754 | -96.8947
24.4792 | -96.8975
24.4831 | -96.9007
24.487 | -96.9047
24.4908 | -96.9086
24.4947 | -96.9155
24.4986 | -96.9232
24.5024 | -96.9321
24.5063 | -96.9428
24.5101 | -96.9558
24.514 | -96.9684
24.5179 | -96.9818
24.5217 | -96.9974
24.5256 | -97.0142
24.5295 | -97.0315
24.5333 | -97.046
24.5372 | -97.0613
24.5411 | -97.0763
24.5449 | -97.0878
24.5488 | -97.0995
24.5527 | -97.1115
24.5565 | -97.123
24.5604 | -97.1331
24.5643 | -97.1436
24.5681 | -97.1546
24.572 | -97.165
24.5758 | -97.1751
24.5797 | -97.1817
24.5836 | -97.1845
24.5874 | -97.18
24.5913 | -97.1697
24.5952 | -97.1611
24.599 | -97.1526
24.6029 | -97.1435
24.6068 | -97.1357
24.6106 | -97.1267
24.6145 | -97.1141
24.6184 | -97.0992
24.6222 | -97.084
24.6261 | -97.0688
24.63 | -97.0524
24.6338 | -97.0356
24.6377 | -97.0166
24.6415 | -96.9949
24.6454 | -96.9699
24.6493 | -96.9432
24.6531 | -96.9136
24.657 | -96.8805
24.6609 | -96.8471
24.6647 | -96.8145
24.6686 | -96.7845
24.6725 | -96.7519
24.6763 | -96.7164
24.6802 | -96.6801
24.6841 | -96.6448
24.6879 | -96.6086
24.6918 | -96.5731
24.6957 | -96.5398
24.6995 | -96.5064
24.7034 | -96.4733
24.7072 | -96.4422
24.7111 | -96.4123
24.715 | -96.3823
24.7188 | -96.3538
24.7227 | -96.3273
24.7266 | -96.3021
24.7304 | -96.2786
24.7343 | -96.2587
24.7382 | -96.2415
24.742 | -96.2266
24.7459 | -96.2131
24.7498 | -96.2014
24.7536 | -96.1946
24.7575 | -96.1907
24.7614 | -96.1882
24.7652 | -96.1897
24.7691 | -96.1947
24.7729 | -96.2026
24.7768 | -96.2139
24.7807 | -96.2281
24.7845 | -96.2439
24.7884 | -96.2638
24.7923 | -96.2864
24.7961 | -96.3118
24.8 | -96.3399
24.8039 | -96.3685
24.8077 | -96.3997
24.8116 | -96.433
24.8155 | -96.4707
24.8193 | -96.5113
24.8232 | -96.5531
24.8271 | -96.5937
24.8309 | -96.6371
24.8348 | -96.6819
24.8386 | -96.7286
24.8425 | -96.7746
24.8464 | -96.82
24.8502 | -96.8644
24.8541 | -96.9078
24.858 | -96.9496
24.8618 | -96.9919
24.8657 | -97.035
24.8696 | -97.0755
24.8734 | -97.1147
24.8773 | -97.1529
24.8812 | -97.1919
24.885 | -97.2295
24.8889 | -97.2641
24.8928 | -97.2967
24.8966 | -97.327
24.9005 | -97.3533
24.9043 | -97.378
24.9082 | -97.4013
24.9121 | -97.4186
24.9159 | -97.4319
24.9198 | -97.4464
24.9237 | -97.4581
24.9275 | -97.4666
24.9314 | -97.4746
24.9353 | -97.4837
24.9391 | -97.4944
24.943 | -97.5034
24.9469 | -97.5083
24.9507 | -97.5062
24.9546 | -97.4992
24.9585 | -97.4888
24.9623 | -97.4782
24.9662 | -97.4694
24.97 | -97.4586
24.9739 | -97.4446
24.9778 | -97.4295
24.9816 | -97.416
24.9855 | -97.4004
24.9894 | -97.3849
24.9932 | -97.37
24.9971 | -97.3548
25.001 | -97.3384
25.0048 | -97.3246
25.0087 | -97.312
25.0126 | -97.2982
25.0164 | -97.284
25.0203 | -97.2705
25.0242 | -97.2607
25.028 | -97.2527
25.0319 | -97.245
25.0357 | -97.2379
25.0396 | -97.2311
25.0435 | -97.2271
25.0473 | -97.2252
25.0512 | -97.2262
25.0551 | -97.2281
25.0589 | -97.2306
25.0628 | -97.2323
25.0667 | -97.2335
25.0705 | -97.2376
25.0744 | -97.2424
25.0783 | -97.2496
25.0821 | -97.2603
25.086 | -97.27
25.0899 | -97.2873
25.0937 | -97.2978
25.0976 | -97.3095
25.1014 | -97.3238
25.1053 | -97.3384
25.1092 | -97.352
25.113 | -97.365
25.1169 | -97.3794
25.1208 | -97.3912
25.1246 | -97.4058
25.1285 | -97.4232
25.1324 | -97.4395
25.1362 | -97.4554
25.1401 | -97.4719
25.144 | -97.4871
25.1478 | -97.5036
25.1517 | -97.5213
25.1556 | -97.5387
25.1594 | -97.5549
25.1633 | -97.5706
25.1671 | -97.5864
25.171 | -97.6011
25.1749 | -97.6137
25.1787 | -97.625
25.1826 | -97.6393
25.1865 | -97.6519
25.1903 | -97.6628
25.1942 | -97.674
25.1981 | -97.6866
25.2019 | -97.6986
25.2058 | -97.7093
25.2097 | -97.7197
25.2135 | -97.7303
25.2174 | -97.7393
25.2213 | -97.7477
25.2251 | -97.7553
25.229 | -97.7617
25.2328 | -97.7656
25.2367 | -97.7684
25.2406 | -97.7706
25.2444 | -97.7722
25.2483 | -97.774
25.2522 | -97.7764
25.256 | -97.7789
25.2599 | -97.7814
25.2638 | -97.7828
25.2676 | -97.7867
25.2715 | -97.7883
25.2754 | -97.7847
25.2792 | -97.7823
25.2831 | -97.78
25.287 | -97.7772
25.2908 | -97.7754
25.2947 | -97.7742
25.2986 | -97.7725
25.3024 | -97.7687
25.3063 | -97.7676
25.3101 | -97.7667
25.314 | -97.7648
25.3179 | -97.7614
25.3217 | -97.7574
25.3256 | -97.756
25.3295 | -97.7562
25.3333 | -97.7573
25.3372 | -97.7577
25.3411 | -97.7605
25.3449 | -97.7649
25.3488 | -97.7676
25.3527 | -97.7715
25.3565 | -97.7768
25.3604 | -97.7826
25.3643 | -97.7895
25.3681 | -97.7965
25.372 | -97.8043
25.3758 | -97.8131
25.3797 | -97.822
25.3836 | -97.8317
25.3874 | -97.8427
25.3913 | -97.8537
25.3952 | -97.8673
25.399 | -97.881
25.4029 | -97.8953
25.4068 | -97.9092
25.4106 | -97.9238
25.4145 | -97.9404
25.4184 | -97.9582
25.4222 | -97.9774
25.4261 | -97.9965
25.43 | -98.0159
25.4338 | -98.0338
25.4377 | -98.0522
25.4415 | -98.0709
25.4454 | -98.0905
25.4493 | -98.1112
25.4531 | -98.1327
25.457 | -98.1528
25.4609 | -98.1704
25.4647 | -98.1883
25.4686 | -98.2046
25.4725 | -98.2222
25.4763 | -98.241
25.4802 | -98.2607
25.4841 | -98.2772
25.4879 | -98.2916
25.4918 | -98.3063
25.4957 | -98.32
25.4995 | -98.3328
25.5034 | -98.346
25.5072 | -98.3567
25.5111 | -98.3651
25.515 | -98.3728
25.5188 | -98.3819
25.5227 | -98.3886
25.5266 | -98.395
25.5304 | -98.4008
25.5343 | -98.4071
25.5382 | -98.4122
25.542 | -98.412
25.5459 | -98.4116
25.5498 | -98.4115
25.5536 | -98.4112
25.5575 | -98.4087
25.5614 | -98.4048
25.5652 | -98.403
25.5691 | -98.4009
25.5729 | -98.3975
25.5768 | -98.3937
25.5807 | -98.3872
25.5845 | -98.3819
25.5884 | -98.3782
25.5923 | -98.3745
25.5961 | -98.3698
25.6 | -98.3662
25.6039 | -98.3626
25.6077 | -98.3575
25.6116 | -98.3509
25.6155 | -98.345
25.6193 | -98.339
25.6232 | -98.3349
25.6271 | -98.3337
25.6309 | -98.3327
25.6348 | -98.3313
25.6386 | -98.3302
25.6425 | -98.3295
25.6464 | -98.3333
25.6502 | -98.3341
25.6541 | -98.3341
25.658 | -98.3362
25.6618 | -98.3392
25.6657 | -98.3437
25.6696 | -98.3489
25.6734 | -98.3551
25.6773 | -98.362
25.6812 | -98.3696
25.685 | -98.3793
25.6889 | -98.3885
25.6928 | -98.3978
25.6966 | -98.4066
25.7005 | -98.4154
25.7043 | -98.4261
25.7082 | -98.4375
25.7121 | -98.452
25.7159 | -98.4663
25.7198 | -98.4798
25.7237 | -98.4942
25.7275 | -98.5097
25.7314 | -98.5266
25.7353 | -98.5426
25.7391 | -98.5584
25.743 | -98.5751
25.7469 | -98.5934
25.7507 | -98.6113
25.7546 | -98.6303
25.7585 | -98.649
25.7623 | -98.6668
25.7662 | -98.6859
25.77 | -98.7055
25.7739 | -98.7244
25.7778 | -98.7447
25.7816 | -98.7647
25.7855 | -98.7859
25.7894 | -98.8059
25.7932 | -98.8275
25.7971 | -98.851
25.801 | -98.8727
25.8048 | -98.8944
25.8087 | -98.9151
25.8126 | -98.9347
25.8164 | -98.9553
25.8203 | -98.9758
25.8242 | -98.9971
25.828 | -99.0189
25.8319 | -99.0415
25.8357 | -99.0627
25.8396 | -99.0823
25.8435 | -99.1026
25.8473 | -99.1225
25.8512 | -99.1402
25.8551 | -99.158
25.8589 | -99.1768
25.8628 | -99.1951
25.8667 | -99.2097
25.8705 | -99.223
25.8744 | -99.2403
25.8783 | -99.2539
25.8821 | -99.2678
25.886 | -99.283
25.8899 | -99.2948
25.8937 | -99.3037
25.8976 | -99.3121
25.9014 | -99.3221
25.9053 | -99.333
25.9092 | -99.3439
25.913 | -99.3548
25.9169 | -99.3614
25.9208 | -99.3673
25.9246 | -99.3752
25.9285 | -99.3807
25.9324 | -99.3844
25.9362 | -99.3879
25.9401 | -99.3889
25.944 | -99.3902
25.9478 | -99.3928
25.9517 | -99.392
25.9556 | -99.3922
25.9594 | -99.3938
25.9633 | -99.3924
25.9672 | -99.3896
25.971 | -99.3874
25.9749 | -99.3875
25.9787 | -99.3866
25.9826 | -99.3837
25.9865 | -99.3775
25.9903 | -99.3697
25.9942 | -99.3613
25.9981 | -99.3557
26.0019 | -99.349
26.0058 | -99.3432
26.0097 | -99.3383
26.0135 | -99.332
26.0174 | -99.3246
26.0213 | -99.3158
26.0251 | -99.3062
26.029 | -99.2978
26.0328 | -99.2888
26.0367 | -99.2778
26.0406 | -99.2667
26.0444 | -99.2572
26.0483 | -99.248
26.0522 | -99.2366
26.056 | -99.2248
26.0599 | -99.212
26.0638 | -99.2
26.0676 | -99.1896
26.0715 | -99.1768
26.0754 | -99.1642
26.0792 | -99.1523
26.0831 | -99.1402
26.087 | -99.1288
26.0908 | -99.1178
26.0947 | -99.1069
26.0986 | -99.0964
26.1024 | -99.0861
26.1063 | -99.0747
26.1101 | -99.062
26.114 | -99.0504
26.1179 | -99.0406
26.1217 | -99.0315
26.1256 | -99.0235
26.1295 | -99.0164
26.1333 | -99.0088
26.1372 | -99.0039
26.1411 | -98.9995
26.1449 | -98.9946
26.1488 | -98.9929
26.1527 | -98.9897
26.1565 | -98.9872
26.1604 | -98.9851
26.1643 | -98.9825
26.1681 | -98.9813
26.172 | -98.9832
26.1758 | -98.9865
26.1797 | -98.989
26.1836 | -98.9923
26.1874 | -98.9985
26.1913 | -99.0051
26.1952 | -99.0117
26.199 | -99.0213
26.2029 | -99.0399
26.2068 | -99.0513
26.2106 | -99.0628
26.2145 | -99.0744
26.2184 | -99.0869
26.2222 | -99.1004
26.2261 | -99.1149
26.23 | -99.1301
26.2338 | -99.1465
26.2377 | -99.1645
26.2415 | -99.185
26.2454 | -99.202
26.2493 | -99.2212
26.2531 | -99.2386
26.257 | -99.2553
26.2609 | -99.275
26.2647 | -99.2958
26.2686 | -99.3169
26.2725 | -99.3364
26.2763 | -99.3572
26.2802 | -99.3779
26.2841 | -99.397
26.2879 | -99.4163
26.2918 | -99.4364
26.2957 | -99.4546
26.2995 | -99.4712
26.3034 | -99.4888
26.3072 | -99.5081
26.3111 | -99.5287
26.315 | -99.5482
26.3188 | -99.5657
26.3227 | -99.5839
26.3266 | -99.601
26.3304 | -99.6163
26.3343 | -99.6332
26.3382 | -99.6516
26.342 | -99.6684
26.3459 | -99.6823
26.3498 | -99.6956
26.3536 | -99.7082
26.3575 | -99.7223
26.3614 | -99.7358
26.3652 | -99.7499
26.3691 | -99.7614
26.3729 | -99.7711
26.3768 | -99.7775
26.3807 | -99.7843
26.3845 | -99.7924
26.3884 | -99.7998
26.3923 | -99.8078
26.3961 | -99.8153
26.4 | -99.8209
26.4039 | -99.8249
26.4077 | -99.8314
26.4116 | -99.8389
26.4155 | -99.8456
26.4193 | -99.8504
26.4232 | -99.8557
26.4271 | -99.8598
26.4309 | -99.8657
26.4348 | -99.8717
26.4386 | -99.8742
26.4425 | -99.8776
26.4464 | -99.8812
26.4502 | -99.8839
26.4541 | -99.8865
26.458 | -99.8908
26.4618 | -99.8938
26.4657 | -99.8964
26.4696 | -99.8986
26.4734 | -99.9008
26.4773 | -99.904
26.4812 | -99.9071
26.485 | -99.9078
26.4889 | -99.9076
26.4928 | -99.9071
26.4966 | -99.9069
26.5005 | -99.9076
26.5043 | -99.9073
26.5082 | -99.9082
26.5121 | -99.9101
26.5159 | -99.9116
26.5198 | -99.9124
26.5237 | -99.9145
26.5275 | -99.9173
26.5314 | -99.9196
26.5353 | -99.9229
26.5391 | -99.9262
26.543 | -99.9297
26.5469 | -99.9331
26.5507 | -99.9365
26.5546 | -99.9414
26.5585 | -99.9454
26.5623 | -99.9502
26.5662 | -99.9566
26.57 | -99.962
26.5739 | -99.9688
26.5778 | -99.9761
26.5816 | -99.9822
26.5855 | -99.9894
26.5894 | -99.9977
26.5932 | -100.0084
26.5971 | -100.0195
26.601 | -100.0295
26.6048 | -100.0396
26.6087 | -100.0504
26.6126 | -100.0641
26.6164 | -100.0761
26.6203 | -100.0911
26.6242 | -100.107
26.628 | -100.1218
26.6319 | -100.1361
26.6357 | -100.1513
26.6396 | -100.1665
26.6435 | -100.1825
26.6473 | -100.1983
26.6512 | -100.2152
26.6551 | -100.2308
26.6589 | -100.2473
26.6628 | -100.2663
26.6667 | -100.2832
26.6705 | -100.302
26.6744 | -100.324
26.6783 | -100.3439
26.6821 | -100.36
26.686 | -100.3768
26.6899 | -100.3935
26.6937 | -100.4101
26.6976 | -100.4261
26.7014 | -100.4424
26.7053 | -100.4584
26.7092 | -100.4738
26.713 | -100.487
26.7169 | -100.5011
26.7208 | -100.5145
26.7246 | -100.5279
26.7285 | -100.5415
26.7324 | -100.5535
26.7362 | -100.5634
26.7401 | -100.5721
26.744 | -100.5823
26.7478 | -100.5927
26.7517 | -100.6045
26.7556 | -100.6142
26.7594 | -100.6167
26.7633 | -100.6237
26.7672 | -100.6308
26.771 | -100.6388
26.7749 | -100.6449
26.7787 | -100.6519
26.7826 | -100.6584
26.7865 | -100.6624
26.7903 | -100.6645
26.7942 | -100.6664
26.7981 | -100.6707
26.8019 | -100.6754
26.8058 | -100.6807
26.8097 | -100.6857
26.8135 | -100.6885
26.8174 | -100.6917
26.8213 | -100.6942
26.8251 | -100.6962
26.829 | -100.6985
26.8329 | -100.7029
26.8367 | -100.7046
26.8406 | -100.7064
26.8444 | -100.7099
26.8483 | -100.7141
26.8522 | -100.7192
26.856 | -100.7242
26.8599 | -100.732
26.8638 | -100.7392
26.8676 | -100.7456
26.8715 | -100.7519
26.8754 | -100.758
26.8792 | -100.7642
26.8831 | -100.7724
26.887 | -100.7794
26.8908 | -100.7862
26.8947 | -100.7943
26.8986 | -100.8021
26.9024 | -100.81
26.9063 | -100.8177
26.9101 | -100.8264
26.914 | -100.8358
26.9179 | -100.8453
26.9217 | -100.855
26.9256 | -100.8656
26.9295 | -100.8775
26.9333 | -100.8878
26.9372 | -100.8968
26.9411 | -100.9066
26.9449 | -100.9176
26.9488 | -100.9303
26.9527 | -100.9428
26.9565 | -100.9545
26.9604 | -100.9692
26.9643 | -100.9837
26.9681 | -100.9951
26.972 | -101.0041
26.9758 | -101.015
26.9797 | -101.0275
26.9836 | -101.0408
26.9874 | -101.0538
26.9913 | -101.0651
26.9952 | -101.0769
26.999 | -101.0861
27.0029 | -101.0954
27.0068 | -101.108
27.0106 | -101.1202
27.0145 | -101.1303
27.0184 | -101.1407
27.0222 | -101.1504
27.0261 | -101.162
27.03 | -101.1728
27.0338 | -101.1846
27.0377 | -101.1963
27.0415 | -101.2075
27.0454 | -101.2185
27.0493 | -101.2281
27.0531 | -101.2354
27.057 | -101.2438
27.0609 | -101.2539
27.0647 | -101.2627
27.0686 | -101.2737
27.0725 | -101.285
27.0763 | -101.2963
27.0802 | -101.306
27.0841 | -101.315
27.0879 | -101.3251
27.0918 | -101.3361
27.0957 | -101.346
27.0995 | -101.358
27.1034 | -101.3678
27.1072 | -101.3767
27.1111 | -101.3861
27.115 | -101.3971
27.1188 | -101.4092
27.1227 | -101.4193
27.1266 | -101.4277
27.1304 | -101.4351
27.1343 | -101.4448
27.1382 | -101.4544
27.142 | -101.4656
27.1459 | -101.4749
27.1498 | -101.4829
27.1536 | -101.4904
27.1575 | -101.5001
27.1614 | -101.5109
27.1652 | -101.5222
27.1691 | -101.534
27.1729 | -101.5454
27.1768 | -101.5557
27.1807 | -101.562
27.1845 | -101.5688
27.1884 | -101.5775
27.1923 | -101.5839
27.1961 | -101.5901
27.2 | -101.5962
27.2039 | -101.6014
27.2077 | -101.6076
27.2116 | -101.6131
27.2155 | -101.6179
27.2193 | -101.6234
27.2232 | -101.6279
27.2271 | -101.6301
27.2309 | -101.6334
27.2348 | -101.637
27.2386 | -101.6397
27.2425 | -101.6391
27.2464 | -101.64
27.2502 | -101.642
27.2541 | -101.6435
27.258 | -101.6426
27.2618 | -101.6407
27.2657 | -101.6386
27.2696 | -101.6362
27.2734 | -101.6322
27.2773 | -101.6272
27.2812 | -101.6221
27.285 | -101.6173
27.2889 | -101.6119
27.2928 | -101.6057
27.2966 | -101.6004
27.3005 | -101.594
27.3043 | -101.5836
27.3082 | -101.5749
27.3121 | -101.5673
27.3159 | -101.5601
27.3198 | -101.548
27.3237 | -101.5415
27.3275 | -101.5339
27.3314 | -101.524
27.3353 | -101.5136
27.3391 | -101.5027
27.343 | -101.4938
27.3469 | -101.4847
27.3507 | -101.4756
27.3546 | -101.4676
27.3585 | -101.4607
27.3623 | -101.4547
27.3662 | -101.4471
27.37 | -101.4375
27.3739 | -101.4294
27.3778 | -101.4219
27.3816 | -101.415
27.3855 | -101.4092
27.3894 | -101.4015
27.3932 | -101.3946
27.3971 | -101.3894
27.401 | -101.3839
27.4048 | -101.3783
27.4087 | -101.3723
27.4126 | -101.3676
27.4164 | -101.3631
27.4203 | -101.3589
27.4242 | -101.3552
27.428 | -101.3521
27.4319 | -101.3508
27.4357 | -101.3492
27.4396 | -101.3486
27.4435 | -101.3494
27.4473 | -101.3481
27.4512 | -101.3482
27.4551 | -101.3481
27.4589 | -101.3448
27.4628 | -101.344
27.4667 | -101.3444
27.4705 | -101.3447
27.4744 | -101.3436
27.4783 | -101.3434
27.4821 | -101.3444
27.486 | -101.3447
27.4899 | -101.347
27.4937 | -101.3468
27.4976 | -101.3468
27.5014 | -101.3479
27.5053 | -101.3489
27.5092 | -101.3486
27.513 | -101.3481
27.5169 | -101.3484
27.5208 | -101.3474
27.5246 | -101.346
27.5285 | -101.3429
27.5324 | -101.3395
27.5362 | -101.3361
27.5401 | -101.3328
27.544 | -101.3291
27.5478 | -101.325
27.5517 | -101.3196
27.5556 | -101.3144
27.5594 | -101.3097
27.5633 | -101.3031
27.5672 | -101.2976
27.571 | -101.2894
27.5749 | -101.2807
27.5787 | -101.2749
27.5826 | -101.2696
27.5865 | -101.2625
27.5903 | -101.2553
27.5942 | -101.246
27.5981 | -101.2373
27.6019 | -101.2294
27.6058 | -101.2212
27.6097 | -101.2118
27.6135 | -101.2024
27.6174 | -101.1889
27.6213 | -101.1775
27.6251 | -101.1654
27.629 | -101.1548
27.6329 | -101.1444
27.6367 | -101.1314
27.6406 | -101.1189
27.6444 | -101.1075
27.6483 | -101.097
27.6522 | -101.0854
27.656 | -101.073
27.6599 | -101.0617
27.6638 | -101.051
27.6676 | -101.0415
27.6715 | -101.0333
27.6754 | -101.0242
27.6792 | -101.0144
27.6831 | -101.0034
27.687 | -100.9936
27.6908 | -100.986
27.6947 | -100.9764
27.6986 | -100.9667
27.7024 | -100.957
27.7063 | -100.9485
27.7101 | -100.9415
27.714 | -100.9332
27.7179 | -100.9241
27.7217 | -100.9159
27.7256 | -100.9086
27.7295 | -100.9027
27.7333 | -100.8951
27.7372 | -100.8863
27.7411 | -100.8781
27.7449 | -100.8706
27.7488 | -100.8625
27.7527 | -100.8526
27.7565 | -100.8436
27.7604 | -100.8368
27.7643 | -100.8283
27.7681 | -100.8222
27.772 | -100.8167
27.7758 | -100.8112
27.7797 | -100.8068
27.7836 | -100.8011
27.7874 | -100.7944
27.7913 | -100.7887
27.7952 | -100.7822
27.799 | -100.7759
27.8029 | -100.7683
27.8068 | -100.7609
27.8106 | -100.7533
27.8145 | -100.7462
27.8184 | -100.7393
27.8222 | -100.732
27.8261 | -100.726
27.83 | -100.7201
27.8338 | -100.7133
27.8377 | -100.7093
27.8415 | -100.7037
27.8454 | -100.6962
27.8493 | -100.6917
27.8531 | -100.6871
27.857 | -100.6817
27.8609 | -100.6778
27.8647 | -100.6755
27.8686 | -100.6717
27.8725 | -100.6672
27.8763 | -100.6653
27.8802 | -100.6612
27.8841 | -100.6559
27.8879 | -100.6513
27.8918 | -100.6479
27.8957 | -100.6446
27.8995 | -100.6416
27.9034 | -100.6393
27.9072 | -100.6389
27.9111 | -100.6383
27.915 | -100.6372
27.9188 | -100.6348
27.9227 | -100.633
27.9266 | -100.6317
27.9304 | -100.6317
27.9343 | -100.6314
27.9382 | -100.6312
27.942 | -100.631
27.9459 | -100.632
27.9498 | -100.6313
27.9536 | -100.6309
27.9575 | -100.6336
27.9614 | -100.6346
27.9652 | -100.6364
27.9691 | -100.6391
27.9729 | -100.6402
27.9768 | -100.6415
27.9807 | -100.6449
27.9845 | -100.649
27.9884 | -100.6533
27.9923 | -100.6573
27.9961 | -100.6601
28 | -100.6641
28.0039 | -100.6684