## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 234 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 236
Hood | 02-92B/P | 175
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1012.1 | hPa |  | PRC PRO 01
Humidity | 46.9 | %/Hr
Internal temperature of calibration bench | 23.99 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1018 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.1 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.05 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.05 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3060 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 23/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 27/05/2018
Engineer | Anthony HUE | Approval |  | 29/05/2018
CEO | Robert GLEMOT | Approval |  | 29/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.348 | 57.5007
0.012209 | -35.912 | 38.3301
0.018313 | -34.8518 | 28.1919
0.024417 | -34.3893 | 21.9243
0.030521 | -34.2786 | 18.1698
0.036626 | -34.1351 | 15.1211
0.04273 | -34.0405 | 12.7582
0.048834 | -33.9955 | 11.2407
0.054938 | -33.9311 | 10.1087
0.061043 | -33.8696 | 9.0251
0.067147 | -33.8591 | 8.1981
0.073251 | -33.8602 | 7.4352
0.079355 | -33.8789 | 6.7735
0.08546 | -33.8931 | 6.1832
0.091564 | -33.8884 | 5.7499
0.097668 | -33.8702 | 5.5063
0.10377 | -33.848 | 5.2545
0.10988 | -33.8308 | 4.9739
0.11598 | -33.8219 | 4.6782
0.12209 | -33.8241 | 4.3936
0.12819 | -33.8345 | 4.1113
0.13429 | -33.8459 | 3.8256
0.1404 | -33.8493 | 3.5493
0.1465 | -33.8513 | 3.2928
0.15261 | -33.8486 | 3.0727
0.15871 | -33.842 | 2.91
0.16482 | -33.8339 | 2.8064
0.17092 | -33.8316 | 2.7672
0.17702 | -33.8294 | 2.6952
0.18313 | -33.8285 | 2.6175
0.18923 | -33.8281 | 2.5094
0.19534 | -33.8267 | 2.3838
0.20144 | -33.8254 | 2.2428
0.20754 | -33.8226 | 2.1228
0.21365 | -33.8216 | 2.0089
0.21975 | -33.8212 | 1.9117
0.22586 | -33.8222 | 1.8249
0.23196 | -33.8247 | 1.7531
0.23807 | -33.8287 | 1.6908
0.24417 | -33.8327 | 1.6326
0.25027 | -33.8367 | 1.5705
0.25638 | -33.839 | 1.5088
0.26248 | -33.8409 | 1.4422
0.26859 | -33.8418 | 1.3768
0.27469 | -33.8423 | 1.3056
0.2808 | -33.8434 | 1.2318
0.2869 | -33.8443 | 1.1608
0.293 | -33.8453 | 1.0937
0.29911 | -33.8463 | 1.0338
0.30521 | -33.8473 | 0.97477
0.31132 | -33.8481 | 0.93645
0.31742 | -33.8494 | 0.89395
0.32353 | -33.8498 | 0.85931
0.32963 | -33.8508 | 0.83071
0.33573 | -33.8519 | 0.81099
0.34184 | -33.8533 | 0.79018
0.34794 | -33.8555 | 0.76705
0.35405 | -33.8577 | 0.74361
0.36015 | -33.8603 | 0.71886
0.36626 | -33.8623 | 0.69196
0.37236 | -33.8645 | 0.65581
0.37846 | -33.8666 | 0.61192
0.38457 | -33.8675 | 0.55503
0.39067 | -33.8685 | 0.50868
0.39678 | -33.8691 | 0.45702
0.40288 | -33.8695 | 0.40315
0.40899 | -33.8693 | 0.35143
0.41509 | -33.8686 | 0.30071
0.42119 | -33.8685 | 0.25748
0.4273 | -33.8684 | 0.21129
0.4334 | -33.8683 | 0.16517
0.43951 | -33.8685 | 0.12176
0.44561 | -33.8693 | 0.082057
0.45172 | -33.871 | 0.050382
0.45782 | -33.8718 | 0.013303
0.46392 | -33.8731 | -0.022808
0.47003 | -33.8742 | -0.055077
0.47613 | -33.8751 | -0.089021
0.48224 | -33.8757 | -0.1237
0.48834 | -33.8748 | -0.15833
0.49445 | -33.8746 | -0.19351
0.50055 | -33.8741 | -0.22545
0.50665 | -33.8734 | -0.25957
0.51276 | -33.8724 | -0.29354
0.51886 | -33.8714 | -0.32741
0.52497 | -33.8701 | -0.35152
0.53107 | -33.8694 | -0.38094
0.53717 | -33.8687 | -0.40911
0.54328 | -33.8678 | -0.43191
0.54938 | -33.8676 | -0.45438
0.55549 | -33.8676 | -0.47526
0.56159 | -33.8687 | -0.4869
0.5677 | -33.869 | -0.50147
0.5738 | -33.8701 | -0.51778
0.5799 | -33.8709 | -0.53272
0.58601 | -33.872 | -0.54283
0.59211 | -33.8737 | -0.54914
0.59822 | -33.875 | -0.55864
0.60432 | -33.8765 | -0.57074
0.61043 | -33.8778 | -0.57869
0.61653 | -33.8791 | -0.58898
0.62263 | -33.8806 | -0.5983
0.62874 | -33.8813 | -0.6089
0.63484 | -33.8822 | -0.61815
0.64095 | -33.8827 | -0.62833
0.64705 | -33.883 | -0.63808
0.65316 | -33.8832 | -0.65046
0.65926 | -33.8833 | -0.66627
0.66536 | -33.8822 | -0.68371
0.67147 | -33.8818 | -0.69689
0.67757 | -33.8814 | -0.71119
0.68368 | -33.8807 | -0.73013
0.68978 | -33.8799 | -0.74735
0.69589 | -33.8789 | -0.76636
0.70199 | -33.878 | -0.78915
0.70809 | -33.8771 | -0.80834
0.7142 | -33.8764 | -0.82881
0.7203 | -33.8753 | -0.84803
0.72641 | -33.8746 | -0.87053
0.73251 | -33.874 | -0.89052
0.73862 | -33.8742 | -0.91221
0.74472 | -33.8739 | -0.93217
0.75082 | -33.8737 | -0.95425
0.75693 | -33.874 | -0.97602
0.76303 | -33.8746 | -0.99793
0.76914 | -33.8765 | -1.0182
0.77524 | -33.8775 | -1.0377
0.78135 | -33.8788 | -1.0577
0.78745 | -33.8802 | -1.0775
0.79355 | -33.8814 | -1.0971
0.79966 | -33.8832 | -1.1163
0.80576 | -33.8855 | -1.1359
0.81187 | -33.8872 | -1.1573
0.81797 | -33.8891 | -1.1785
0.82408 | -33.8909 | -1.1989
0.83018 | -33.8926 | -1.2193
0.83628 | -33.8945 | -1.2377
0.84239 | -33.8957 | -1.2583
0.84849 | -33.8972 | -1.28
0.8546 | -33.8987 | -1.2998
0.8607 | -33.8999 | -1.3205
0.86681 | -33.901 | -1.3394
0.87291 | -33.902 | -1.3583
0.87901 | -33.9025 | -1.3764
0.88512 | -33.9033 | -1.3946
0.89122 | -33.904 | -1.4129
0.89733 | -33.9044 | -1.4285
0.90343 | -33.9048 | -1.4462
0.90953 | -33.905 | -1.4606
0.91564 | -33.905 | -1.4764
0.92174 | -33.9053 | -1.4914
0.92785 | -33.9054 | -1.5061
0.93395 | -33.9055 | -1.5199
0.94006 | -33.9056 | -1.5327
0.94616 | -33.9054 | -1.5445
0.95226 | -33.9055 | -1.5571
0.95837 | -33.9056 | -1.5709
0.96447 | -33.9058 | -1.5836
0.97058 | -33.9057 | -1.597
0.97668 | -33.9055 | -1.6102
0.98279 | -33.9052 | -1.6228
0.98889 | -33.905 | -1.6334
0.99499 | -33.905 | -1.6466
1.0011 | -33.9049 | -1.6572
1.0072 | -33.9045 | -1.6686
1.0133 | -33.9042 | -1.6792
1.0194 | -33.9038 | -1.692
1.0255 | -33.9034 | -1.7056
1.0316 | -33.9029 | -1.7173
1.0377 | -33.9025 | -1.7308
1.0438 | -33.9019 | -1.7434
1.0499 | -33.9011 | -1.7585
1.056 | -33.9004 | -1.7727
1.0621 | -33.8998 | -1.7868
1.0682 | -33.8993 | -1.7991
1.0743 | -33.8987 | -1.8144
1.0805 | -33.8982 | -1.8286
1.0866 | -33.8972 | -1.8435
1.0927 | -33.8964 | -1.8578
1.0988 | -33.8957 | -1.8726
1.1049 | -33.8949 | -1.8866
1.111 | -33.8941 | -1.9018
1.1171 | -33.8934 | -1.9162
1.1232 | -33.8929 | -1.9313
1.1293 | -33.8925 | -1.9451
1.1354 | -33.8922 | -1.9599
1.1415 | -33.8918 | -1.9753
1.1476 | -33.8912 | -1.9907
1.1537 | -33.8907 | -2.0052
1.1598 | -33.8908 | -2.0198
1.1659 | -33.8905 | -2.0329
1.172 | -33.8903 | -2.0467
1.1781 | -33.89 | -2.0605
1.1842 | -33.8899 | -2.0745
1.1903 | -33.8903 | -2.0853
1.1964 | -33.8902 | -2.0986
1.2025 | -33.8903 | -2.1124
1.2086 | -33.8906 | -2.1246
1.2147 | -33.8906 | -2.1369
1.2209 | -33.891 | -2.1479
1.227 | -33.8918 | -2.1585
1.2331 | -33.8922 | -2.1703
1.2392 | -33.8925 | -2.182
1.2453 | -33.893 | -2.1929
1.2514 | -33.8934 | -2.2026
1.2575 | -33.894 | -2.2134
1.2636 | -33.8949 | -2.2222
1.2697 | -33.8955 | -2.2313
1.2758 | -33.8962 | -2.2422
1.2819 | -33.8969 | -2.252
1.288 | -33.8976 | -2.2626
1.2941 | -33.8985 | -2.2709
1.3002 | -33.8992 | -2.2788
1.3063 | -33.8997 | -2.2877
1.3124 | -33.9004 | -2.2965
1.3185 | -33.9011 | -2.3054
1.3246 | -33.9018 | -2.3144
1.3307 | -33.9027 | -2.3228
1.3368 | -33.9035 | -2.3313
1.3429 | -33.9042 | -2.3401
1.349 | -33.9047 | -2.3493
1.3551 | -33.9054 | -2.3585
1.3613 | -33.906 | -2.3666
1.3674 | -33.9066 | -2.3766
1.3735 | -33.9071 | -2.3862
1.3796 | -33.9075 | -2.3951
1.3857 | -33.9079 | -2.4042
1.3918 | -33.9084 | -2.4142
1.3979 | -33.909 | -2.4232
1.404 | -33.9095 | -2.4344
1.4101 | -33.9099 | -2.4446
1.4162 | -33.9103 | -2.4551
1.4223 | -33.9109 | -2.4663
1.4284 | -33.9113 | -2.4762
1.4345 | -33.9116 | -2.4876
1.4406 | -33.9119 | -2.5009
1.4467 | -33.9124 | -2.5124
1.4528 | -33.9128 | -2.5236
1.4589 | -33.9133 | -2.5353
1.465 | -33.9135 | -2.546
1.4711 | -33.914 | -2.5572
1.4772 | -33.9146 | -2.5704
1.4833 | -33.915 | -2.5832
1.4894 | -33.9156 | -2.5953
1.4955 | -33.9161 | -2.6071
1.5016 | -33.9166 | -2.6194
1.5078 | -33.917 | -2.6332
1.5139 | -33.9174 | -2.646
1.52 | -33.9179 | -2.6585
1.5261 | -33.9182 | -2.6709
1.5322 | -33.9185 | -2.6848
1.5383 | -33.9189 | -2.6972
1.5444 | -33.9194 | -2.7097
1.5505 | -33.9199 | -2.7229
1.5566 | -33.9204 | -2.7358
1.5627 | -33.9209 | -2.7494
1.5688 | -33.9215 | -2.7628
1.5749 | -33.9219 | -2.7753
1.581 | -33.9221 | -2.7884
1.5871 | -33.9225 | -2.802
1.5932 | -33.9229 | -2.8147
1.5993 | -33.9233 | -2.8288
1.6054 | -33.9237 | -2.8419
1.6115 | -33.9241 | -2.8558
1.6176 | -33.9244 | -2.8682
1.6237 | -33.9246 | -2.8817
1.6298 | -33.9248 | -2.8941
1.6359 | -33.9252 | -2.9066
1.642 | -33.9256 | -2.9184
1.6482 | -33.9258 | -2.9312
1.6543 | -33.9262 | -2.9447
1.6604 | -33.9265 | -2.9565
1.6665 | -33.9266 | -2.9688
1.6726 | -33.9268 | -2.9812
1.6787 | -33.9271 | -2.9935
1.6848 | -33.927 | -3.0059
1.6909 | -33.9273 | -3.0184
1.697 | -33.9275 | -3.0305
1.7031 | -33.9275 | -3.0429
1.7092 | -33.9275 | -3.0548
1.7153 | -33.9277 | -3.0666
1.7214 | -33.9274 | -3.0787
1.7275 | -33.9274 | -3.0911
1.7336 | -33.9273 | -3.1039
1.7397 | -33.9275 | -3.1161
1.7458 | -33.9274 | -3.1281
1.7519 | -33.9275 | -3.1402
1.758 | -33.9275 | -3.1519
1.7641 | -33.9276 | -3.1634
1.7702 | -33.9276 | -3.1762
1.7763 | -33.9275 | -3.1875
1.7824 | -33.9275 | -3.1998
1.7885 | -33.9275 | -3.2122
1.7947 | -33.9276 | -3.224
1.8008 | -33.9276 | -3.2351
1.8069 | -33.9275 | -3.2465
1.813 | -33.9276 | -3.2585
1.8191 | -33.9276 | -3.2711
1.8252 | -33.9277 | -3.2838
1.8313 | -33.9278 | -3.2959
1.8374 | -33.9279 | -3.3076
1.8435 | -33.928 | -3.3194
1.8496 | -33.928 | -3.3324
1.8557 | -33.9282 | -3.345
1.8618 | -33.9284 | -3.3572
1.8679 | -33.9285 | -3.3691
1.874 | -33.9286 | -3.381
1.8801 | -33.9288 | -3.3933
1.8862 | -33.929 | -3.4051
1.8923 | -33.9291 | -3.4172
1.8984 | -33.9292 | -3.4292
1.9045 | -33.9293 | -3.4407
1.9106 | -33.9294 | -3.4528
1.9167 | -33.9294 | -3.4649
1.9228 | -33.9295 | -3.4767
1.9289 | -33.9296 | -3.4885
1.9351 | -33.9298 | -3.4993
1.9412 | -33.93 | -3.5105
1.9473 | -33.9303 | -3.5224
1.9534 | -33.9305 | -3.5338
1.9595 | -33.9307 | -3.5455
1.9656 | -33.9309 | -3.5578
1.9717 | -33.9311 | -3.5694
1.9778 | -33.9315 | -3.5807
1.9839 | -33.9317 | -3.5922
1.99 | -33.9319 | -3.6032
1.9961 | -33.932 | -3.6153
2.0022 | -33.9322 | -3.6266
2.0083 | -33.9323 | -3.6378
2.0144 | -33.9325 | -3.6491
2.0205 | -33.9328 | -3.6608
2.0266 | -33.9331 | -3.6721
2.0327 | -33.9333 | -3.6836
2.0388 | -33.9338 | -3.6952
2.0449 | -33.9339 | -3.7065
2.051 | -33.9342 | -3.7178
2.0571 | -33.9345 | -3.7291
2.0632 | -33.9347 | -3.7405
2.0693 | -33.9349 | -3.7526
2.0754 | -33.9352 | -3.7639
2.0816 | -33.9355 | -3.776
2.0877 | -33.9357 | -3.7874
2.0938 | -33.9361 | -3.7995
2.0999 | -33.9362 | -3.8111
2.106 | -33.9364 | -3.8227
2.1121 | -33.9366 | -3.8346
2.1182 | -33.9369 | -3.8453
2.1243 | -33.9371 | -3.8559
2.1304 | -33.9373 | -3.867
2.1365 | -33.9376 | -3.8787
2.1426 | -33.9379 | -3.8906
2.1487 | -33.9382 | -3.9017
2.1548 | -33.9386 | -3.9135
2.1609 | -33.939 | -3.9248
2.167 | -33.9394 | -3.9359
2.1731 | -33.9397 | -3.9479
2.1792 | -33.9399 | -3.9588
2.1853 | -33.9403 | -3.9707
2.1914 | -33.9406 | -3.982
2.1975 | -33.941 | -3.9934
2.2036 | -33.9412 | -4.0055
2.2097 | -33.9415 | -4.0168
2.2158 | -33.9419 | -4.0274
2.222 | -33.9421 | -4.0377
2.2281 | -33.9424 | -4.0491
2.2342 | -33.9427 | -4.0603
2.2403 | -33.9429 | -4.0708
2.2464 | -33.9432 | -4.0814
2.2525 | -33.9434 | -4.0927
2.2586 | -33.9438 | -4.1032
2.2647 | -33.9441 | -4.1135
2.2708 | -33.9444 | -4.1246
2.2769 | -33.9448 | -4.1352
2.283 | -33.9449 | -4.1462
2.2891 | -33.9452 | -4.157
2.2952 | -33.9454 | -4.1674
2.3013 | -33.9456 | -4.1782
2.3074 | -33.9459 | -4.1885
2.3135 | -33.9461 | -4.1998
2.3196 | -33.9463 | -4.2099
2.3257 | -33.9465 | -4.2204
2.3318 | -33.9467 | -4.2314
2.3379 | -33.9469 | -4.2421
2.344 | -33.947 | -4.2533
2.3501 | -33.9472 | -4.2639
2.3562 | -33.9474 | -4.2742
2.3623 | -33.9476 | -4.2844
2.3685 | -33.9477 | -4.2948
2.3746 | -33.9478 | -4.3059
2.3807 | -33.948 | -4.3163
2.3868 | -33.9482 | -4.3269
2.3929 | -33.9484 | -4.3374
2.399 | -33.9485 | -4.3478
2.4051 | -33.9488 | -4.3581
2.4112 | -33.949 | -4.3686
2.4173 | -33.9492 | -4.3791
2.4234 | -33.9492 | -4.3898
2.4295 | -33.9494 | -4.4005
2.4356 | -33.9496 | -4.4112
2.4417 | -33.9497 | -4.4222
2.4478 | -33.9498 | -4.4329
2.4539 | -33.95 | -4.443
2.46 | -33.9501 | -4.4543
2.4661 | -33.9503 | -4.4653
2.4722 | -33.9504 | -4.4753
2.4783 | -33.9505 | -4.4856
2.4844 | -33.9506 | -4.4965
2.4905 | -33.9507 | -4.5073
2.4966 | -33.9509 | -4.518
2.5027 | -33.951 | -4.5284
2.5089 | -33.9511 | -4.5391
2.515 | -33.9511 | -4.5498
2.5211 | -33.9512 | -4.5604
2.5272 | -33.9513 | -4.5712
2.5333 | -33.9514 | -4.582
2.5394 | -33.9516 | -4.5926
2.5455 | -33.9517 | -4.6039
2.5516 | -33.9519 | -4.6141
2.5577 | -33.952 | -4.6254
2.5638 | -33.9522 | -4.637
2.5699 | -33.9522 | -4.6483
2.576 | -33.9523 | -4.6597
2.5821 | -33.9524 | -4.6708
2.5882 | -33.9525 | -4.6818
2.5943 | -33.9526 | -4.6926
2.6004 | -33.9528 | -4.7034
2.6065 | -33.9529 | -4.7137
2.6126 | -33.9531 | -4.724
2.6187 | -33.9532 | -4.7343
2.6248 | -33.9533 | -4.745
2.6309 | -33.9535 | -4.7563
2.637 | -33.9537 | -4.7678
2.6431 | -33.9539 | -4.7781
2.6492 | -33.954 | -4.7893
2.6554 | -33.9542 | -4.8001
2.6615 | -33.9543 | -4.8109
2.6676 | -33.9545 | -4.8222
2.6737 | -33.9548 | -4.8333
2.6798 | -33.955 | -4.8443
2.6859 | -33.9552 | -4.8554
2.692 | -33.9554 | -4.8662
2.6981 | -33.9556 | -4.8772
2.7042 | -33.9558 | -4.8882
2.7103 | -33.9559 | -4.8988
2.7164 | -33.9562 | -4.9096
2.7225 | -33.9565 | -4.9209
2.7286 | -33.9567 | -4.9317
2.7347 | -33.9569 | -4.9421
2.7408 | -33.9572 | -4.9521
2.7469 | -33.9574 | -4.9633
2.753 | -33.9577 | -4.9738
2.7591 | -33.958 | -4.985
2.7652 | -33.9582 | -4.9963
2.7713 | -33.9583 | -5.007
2.7774 | -33.9586 | -5.0178
2.7835 | -33.9588 | -5.0287
2.7896 | -33.959 | -5.0391
2.7958 | -33.9592 | -5.0497
2.8019 | -33.9595 | -5.06
2.808 | -33.9598 | -5.0707
2.8141 | -33.9601 | -5.0817
2.8202 | -33.9604 | -5.0925
2.8263 | -33.9607 | -5.1033
2.8324 | -33.961 | -5.1145
2.8385 | -33.9614 | -5.1253
2.8446 | -33.9618 | -5.1359
2.8507 | -33.9621 | -5.1466
2.8568 | -33.9625 | -5.157
2.8629 | -33.9627 | -5.1675
2.869 | -33.963 | -5.1785
2.8751 | -33.9634 | -5.1891
2.8812 | -33.9638 | -5.2003
2.8873 | -33.9641 | -5.2111
2.8934 | -33.9644 | -5.222
2.8995 | -33.9647 | -5.233
2.9056 | -33.965 | -5.2441
2.9117 | -33.9653 | -5.255
2.9178 | -33.9657 | -5.2656
2.9239 | -33.9661 | -5.2771
2.93 | -33.9665 | -5.2884
2.9361 | -33.9668 | -5.2991
2.9423 | -33.9671 | -5.3097
2.9484 | -33.9675 | -5.3207
2.9545 | -33.9679 | -5.3316
2.9606 | -33.9683 | -5.3429
2.9667 | -33.9686 | -5.3539
2.9728 | -33.969 | -5.3651
2.9789 | -33.9694 | -5.3761
2.985 | -33.9698 | -5.3873
2.9911 | -33.9702 | -5.3985
2.9972 | -33.9707 | -5.4095
3.0033 | -33.9711 | -5.4209
3.0094 | -33.9715 | -5.4319
3.0155 | -33.9719 | -5.4428
3.0216 | -33.9723 | -5.4539
3.0277 | -33.9726 | -5.4655
3.0338 | -33.973 | -5.4768
3.0399 | -33.9734 | -5.488
3.046 | -33.9737 | -5.4995
3.0521 | -33.974 | -5.5106
3.0582 | -33.9743 | -5.5226
3.0643 | -33.9747 | -5.534
3.0704 | -33.975 | -5.5457
3.0765 | -33.9753 | -5.5567
3.0827 | -33.9756 | -5.5679
3.0888 | -33.9759 | -5.5791
3.0949 | -33.9763 | -5.5903
3.101 | -33.9767 | -5.6017
3.1071 | -33.977 | -5.6134
3.1132 | -33.9774 | -5.6251
3.1193 | -33.9777 | -5.6365
3.1254 | -33.9781 | -5.6478
3.1315 | -33.9785 | -5.6596
3.1376 | -33.9788 | -5.671
3.1437 | -33.9791 | -5.6824
3.1498 | -33.9795 | -5.6944
3.1559 | -33.9798 | -5.7058
3.162 | -33.9801 | -5.7178
3.1681 | -33.9803 | -5.7297
3.1742 | -33.9807 | -5.7412
3.1803 | -33.981 | -5.7528
3.1864 | -33.9814 | -5.764
3.1925 | -33.9817 | -5.7755
3.1986 | -33.9819 | -5.7874
3.2047 | -33.9822 | -5.7994
3.2108 | -33.9825 | -5.8105
3.2169 | -33.9828 | -5.8222
3.223 | -33.983 | -5.8338
3.2292 | -33.9833 | -5.8456
3.2353 | -33.9835 | -5.8564
3.2414 | -33.9837 | -5.8677
3.2475 | -33.9839 | -5.8791
3.2536 | -33.9843 | -5.8905
3.2597 | -33.9845 | -5.9014
3.2658 | -33.9847 | -5.9122
3.2719 | -33.9849 | -5.9232
3.278 | -33.9852 | -5.9347
3.2841 | -33.9855 | -5.9458
3.2902 | -33.9857 | -5.9572
3.2963 | -33.986 | -5.9686
3.3024 | -33.9861 | -5.9796
3.3085 | -33.9864 | -5.9911
3.3146 | -33.9865 | -6.0022
3.3207 | -33.9868 | -6.0136
3.3268 | -33.987 | -6.025
3.3329 | -33.9872 | -6.0357
3.339 | -33.9874 | -6.0465
3.3451 | -33.9876 | -6.0577
3.3512 | -33.9877 | -6.0683
3.3573 | -33.988 | -6.0795
3.3634 | -33.9881 | -6.0906
3.3696 | -33.9883 | -6.1019
3.3757 | -33.9885 | -6.1126
3.3818 | -33.9887 | -6.1234
3.3879 | -33.9889 | -6.1344
3.394 | -33.989 | -6.1457
3.4001 | -33.9892 | -6.1564
3.4062 | -33.9893 | -6.1673
3.4123 | -33.9895 | -6.1779
3.4184 | -33.9897 | -6.1891
3.4245 | -33.9899 | -6.1998
3.4306 | -33.99 | -6.2104
3.4367 | -33.9902 | -6.2209
3.4428 | -33.9903 | -6.2312
3.4489 | -33.9905 | -6.2421
3.455 | -33.9906 | -6.2529
3.4611 | -33.9908 | -6.264
3.4672 | -33.9909 | -6.275
3.4733 | -33.9911 | -6.286
3.4794 | -33.9913 | -6.2968
3.4855 | -33.9914 | -6.3071
3.4916 | -33.9916 | -6.3174
3.4977 | -33.9917 | -6.3286
3.5038 | -33.9919 | -6.3394
3.5099 | -33.992 | -6.3496
3.5161 | -33.9922 | -6.3599
3.5222 | -33.9923 | -6.3702
3.5283 | -33.9924 | -6.3809
3.5344 | -33.9925 | -6.3916
3.5405 | -33.9927 | -6.4021
3.5466 | -33.9928 | -6.4129
3.5527 | -33.9929 | -6.4236
3.5588 | -33.993 | -6.4341
3.5649 | -33.9931 | -6.4447
3.571 | -33.9932 | -6.4553
3.5771 | -33.9933 | -6.4663
3.5832 | -33.9935 | -6.4761
3.5893 | -33.9937 | -6.4866
3.5954 | -33.9939 | -6.4974
3.6015 | -33.994 | -6.5082
3.6076 | -33.9943 | -6.5186
3.6137 | -33.9944 | -6.5287
3.6198 | -33.9947 | -6.5388
3.6259 | -33.9948 | -6.5496
3.632 | -33.995 | -6.5602
3.6381 | -33.9951 | -6.5708
3.6442 | -33.9953 | -6.5813
3.6503 | -33.9955 | -6.5914
3.6565 | -33.9957 | -6.6021
3.6626 | -33.9958 | -6.6126
3.6687 | -33.996 | -6.6232
3.6748 | -33.9963 | -6.6335
3.6809 | -33.9965 | -6.6441
3.687 | -33.9967 | -6.655
3.6931 | -33.9969 | -6.6651
3.6992 | -33.9971 | -6.6756
3.7053 | -33.9973 | -6.686
3.7114 | -33.9975 | -6.6966
3.7175 | -33.9977 | -6.7072
3.7236 | -33.9979 | -6.7171
3.7297 | -33.9981 | -6.7278
3.7358 | -33.9983 | -6.7387
3.7419 | -33.9985 | -6.7494
3.748 | -33.9987 | -6.7599
3.7541 | -33.9989 | -6.7704
3.7602 | -33.9992 | -6.7808
3.7663 | -33.9994 | -6.7911
3.7724 | -33.9997 | -6.8017
3.7785 | -34 | -6.812
3.7846 | -34.0001 | -6.8224
3.7907 | -34.0004 | -6.8332
3.7969 | -34.0007 | -6.8438
3.803 | -34.001 | -6.854
3.8091 | -34.0012 | -6.8644
3.8152 | -34.0015 | -6.8748
3.8213 | -34.0017 | -6.8855
3.8274 | -34.0019 | -6.8958
3.8335 | -34.0023 | -6.9064
3.8396 | -34.0026 | -6.917
3.8457 | -34.0029 | -6.9276
3.8518 | -34.0032 | -6.938
3.8579 | -34.0035 | -6.9485
3.864 | -34.0038 | -6.9592
3.8701 | -34.0041 | -6.9701
3.8762 | -34.0044 | -6.9805
3.8823 | -34.0046 | -6.9912
3.8884 | -34.0049 | -7.0018
3.8945 | -34.0053 | -7.0126
3.9006 | -34.0056 | -7.0235
3.9067 | -34.0059 | -7.0345
3.9128 | -34.0062 | -7.0452
3.9189 | -34.0065 | -7.0555
3.925 | -34.0067 | -7.0662
3.9311 | -34.007 | -7.0764
3.9372 | -34.0073 | -7.0872
3.9434 | -34.0076 | -7.0978
3.9495 | -34.008 | -7.1085
3.9556 | -34.0083 | -7.1189
3.9617 | -34.0086 | -7.1292
3.9678 | -34.0089 | -7.1395
3.9739 | -34.0092 | -7.15
3.98 | -34.0095 | -7.1608
3.9861 | -34.0098 | -7.1713
3.9922 | -34.01 | -7.1819
3.9983 | -34.0103 | -7.1921
4.0044 | -34.0106 | -7.2022
4.0105 | -34.0109 | -7.2128
4.0166 | -34.0112 | -7.2235
4.0227 | -34.0115 | -7.234
4.0288 | -34.0118 | -7.2451
4.0349 | -34.012 | -7.2554
4.041 | -34.0123 | -7.2659
4.0471 | -34.0125 | -7.2766
4.0532 | -34.0128 | -7.2868
4.0593 | -34.0132 | -7.2976
4.0654 | -34.0135 | -7.3079
4.0715 | -34.0138 | -7.319
4.0776 | -34.0141 | -7.3294
4.0838 | -34.0144 | -7.3404
4.0899 | -34.0147 | -7.3507
4.096 | -34.015 | -7.3612
4.1021 | -34.0152 | -7.3715
4.1082 | -34.0155 | -7.3818
4.1143 | -34.0157 | -7.392
4.1204 | -34.0161 | -7.4023
4.1265 | -34.0164 | -7.4125
4.1326 | -34.0166 | -7.4227
4.1387 | -34.017 | -7.433
4.1448 | -34.0172 | -7.4436
4.1509 | -34.0175 | -7.4541
4.157 | -34.0177 | -7.465
4.1631 | -34.018 | -7.4757
4.1692 | -34.0183 | -7.486
4.1753 | -34.0186 | -7.4964
4.1814 | -34.0188 | -7.5065
4.1875 | -34.0191 | -7.5173
4.1936 | -34.0193 | -7.5279
4.1997 | -34.0196 | -7.538
4.2058 | -34.0198 | -7.5487
4.2119 | -34.0201 | -7.559
4.218 | -34.0203 | -7.5692
4.2241 | -34.0206 | -7.5794
4.2303 | -34.0209 | -7.5897
4.2364 | -34.0211 | -7.6001
4.2425 | -34.0213 | -7.6105
4.2486 | -34.0215 | -7.6211
4.2547 | -34.0217 | -7.6312
4.2608 | -34.0219 | -7.6417
4.2669 | -34.0221 | -7.6518
4.273 | -34.0223 | -7.6623
4.2791 | -34.0226 | -7.6727
4.2852 | -34.0228 | -7.6832
4.2913 | -34.023 | -7.6933
4.2974 | -34.0232 | -7.7039
4.3035 | -34.0235 | -7.7146
4.3096 | -34.0237 | -7.7252
4.3157 | -34.0239 | -7.7353
4.3218 | -34.0241 | -7.7456
4.3279 | -34.0244 | -7.756
4.334 | -34.0245 | -7.7663
4.3401 | -34.0248 | -7.7762
4.3462 | -34.025 | -7.7866
4.3523 | -34.0253 | -7.7965
4.3584 | -34.0255 | -7.8064
4.3645 | -34.0257 | -7.8166
4.3707 | -34.0259 | -7.8266
4.3768 | -34.0261 | -7.8369
4.3829 | -34.0263 | -7.8475
4.389 | -34.0265 | -7.8578
4.3951 | -34.0268 | -7.8679
4.4012 | -34.0269 | -7.8785
4.4073 | -34.0271 | -7.8885
4.4134 | -34.0273 | -7.8987
4.4195 | -34.0276 | -7.9089
4.4256 | -34.0279 | -7.9189
4.4317 | -34.0281 | -7.9291
4.4378 | -34.0283 | -7.939
4.4439 | -34.0285 | -7.9489
4.45 | -34.0286 | -7.9591
4.4561 | -34.0288 | -7.9693
4.4622 | -34.029 | -7.9796
4.4683 | -34.0292 | -7.9897
4.4744 | -34.0294 | -8.0002
4.4805 | -34.0296 | -8.01
4.4866 | -34.0298 | -8.0202
4.4927 | -34.03 | -8.0306
4.4988 | -34.0302 | -8.0408
4.5049 | -34.0305 | -8.0506
4.511 | -34.0307 | -8.0604
4.5172 | -34.0309 | -8.0705
4.5233 | -34.0311 | -8.0811
4.5294 | -34.0313 | -8.0911
4.5355 | -34.0315 | -8.1013
4.5416 | -34.0317 | -8.1114
4.5477 | -34.032 | -8.1215
4.5538 | -34.0321 | -8.1312
4.5599 | -34.0324 | -8.1413
4.566 | -34.0326 | -8.1514
4.5721 | -34.0328 | -8.1616
4.5782 | -34.033 | -8.172
4.5843 | -34.0332 | -8.1822
4.5904 | -34.0335 | -8.1923
4.5965 | -34.0337 | -8.2024
4.6026 | -34.034 | -8.213
4.6087 | -34.0341 | -8.2229
4.6148 | -34.0343 | -8.2334
4.6209 | -34.0346 | -8.2435
4.627 | -34.0348 | -8.2539
4.6331 | -34.035 | -8.2638
4.6392 | -34.0352 | -8.2741
4.6453 | -34.0354 | -8.2842
4.6514 | -34.0356 | -8.2944
4.6576 | -34.0358 | -8.3049
4.6637 | -34.036 | -8.3147
4.6698 | -34.0362 | -8.3247
4.6759 | -34.0364 | -8.3352
4.682 | -34.0366 | -8.3452
4.6881 | -34.0367 | -8.3553
4.6942 | -34.037 | -8.3654
4.7003 | -34.0372 | -8.3756
4.7064 | -34.0374 | -8.3855
4.7125 | -34.0377 | -8.396
4.7186 | -34.0378 | -8.4058
4.7247 | -34.0381 | -8.4161
4.7308 | -34.0383 | -8.4261
4.7369 | -34.0385 | -8.4364
4.743 | -34.0388 | -8.4469
4.7491 | -34.039 | -8.4574
4.7552 | -34.0392 | -8.4677
4.7613 | -34.0393 | -8.4779
4.7674 | -34.0395 | -8.4879
4.7735 | -34.0397 | -8.4986
4.7796 | -34.04 | -8.509
4.7857 | -34.0402 | -8.5193
4.7918 | -34.0405 | -8.5293
4.7979 | -34.0407 | -8.5394
4.8041 | -34.0409 | -8.5496
4.8102 | -34.0412 | -8.5595
4.8163 | -34.0414 | -8.5696
4.8224 | -34.0416 | -8.5798
4.8285 | -34.0418 | -8.5897
4.8346 | -34.0421 | -8.5997
4.8407 | -34.0423 | -8.6099
4.8468 | -34.0425 | -8.6204
4.8529 | -34.0428 | -8.6308
4.859 | -34.043 | -8.6414
4.8651 | -34.0432 | -8.6513
4.8712 | -34.0435 | -8.6613
4.8773 | -34.0437 | -8.6716
4.8834 | -34.0439 | -8.6818
4.8895 | -34.0441 | -8.6918
4.8956 | -34.0443 | -8.702
4.9017 | -34.0445 | -8.712
4.9078 | -34.0447 | -8.7225
4.9139 | -34.045 | -8.7322
4.92 | -34.0453 | -8.7421
4.9261 | -34.0455 | -8.752
4.9322 | -34.0458 | -8.7622
4.9383 | -34.046 | -8.7724
4.9445 | -34.0463 | -8.7825
4.9506 | -34.0465 | -8.7925
4.9567 | -34.0467 | -8.8025
4.9628 | -34.047 | -8.8126
4.9689 | -34.0472 | -8.8231
4.975 | -34.0474 | -8.8336
4.9811 | -34.0477 | -8.8435
4.9872 | -34.048 | -8.8536
4.9933 | -34.0483 | -8.864
4.9994 | -34.0485 | -8.8746
5.0055 | -34.0488 | -8.8849
5.0116 | -34.049 | -8.8955
5.0177 | -34.0493 | -8.906
5.0238 | -34.0496 | -8.9164
5.0299 | -34.0499 | -8.927
5.036 | -34.0502 | -8.9373
5.0421 | -34.0504 | -8.9474
5.0482 | -34.0507 | -8.9576
5.0543 | -34.0509 | -8.9676
5.0604 | -34.0512 | -8.9778
5.0665 | -34.0515 | -8.9878
5.0726 | -34.0518 | -8.9983
5.0787 | -34.0521 | -9.0088
5.0848 | -34.0523 | -9.0191
5.091 | -34.0526 | -9.0291
5.0971 | -34.0529 | -9.039
5.1032 | -34.0531 | -9.0496
5.1093 | -34.0534 | -9.0598
5.1154 | -34.0536 | -9.0698
5.1215 | -34.0539 | -9.0798
5.1276 | -34.0542 | -9.0899
5.1337 | -34.0545 | -9.0998
5.1398 | -34.0547 | -9.1098
5.1459 | -34.055 | -9.1203
5.152 | -34.0553 | -9.1307
5.1581 | -34.0555 | -9.1411
5.1642 | -34.0558 | -9.1514
5.1703 | -34.0561 | -9.1617
5.1764 | -34.0564 | -9.1718
5.1825 | -34.0567 | -9.1817
5.1886 | -34.057 | -9.1917
5.1947 | -34.0572 | -9.2019
5.2008 | -34.0575 | -9.2123
5.2069 | -34.0578 | -9.2228
5.213 | -34.0581 | -9.2328
5.2191 | -34.0584 | -9.2431
5.2252 | -34.0586 | -9.2534
5.2314 | -34.0589 | -9.2638
5.2375 | -34.0592 | -9.2738
5.2436 | -34.0594 | -9.2842
5.2497 | -34.0597 | -9.2941
5.2558 | -34.06 | -9.3047
5.2619 | -34.0603 | -9.3153
5.268 | -34.0606 | -9.3257
5.2741 | -34.0609 | -9.336
5.2802 | -34.0612 | -9.3467
5.2863 | -34.0615 | -9.3569
5.2924 | -34.0618 | -9.3674
5.2985 | -34.0622 | -9.3777
5.3046 | -34.0624 | -9.3881
5.3107 | -34.0627 | -9.3986
5.3168 | -34.063 | -9.409
5.3229 | -34.0633 | -9.4193
5.329 | -34.0636 | -9.4296
5.3351 | -34.0639 | -9.4396
5.3412 | -34.0642 | -9.4499
5.3473 | -34.0645 | -9.4603
5.3534 | -34.0648 | -9.4707
5.3595 | -34.0651 | -9.4811
5.3656 | -34.0654 | -9.4914
5.3717 | -34.0657 | -9.5021
5.3779 | -34.066 | -9.5127
5.384 | -34.0663 | -9.523
5.3901 | -34.0666 | -9.5337
5.3962 | -34.0668 | -9.5438
5.4023 | -34.0672 | -9.5544
5.4084 | -34.0675 | -9.5647
5.4145 | -34.0678 | -9.5748
5.4206 | -34.0681 | -9.5851
5.4267 | -34.0684 | -9.5958
5.4328 | -34.0687 | -9.6062
5.4389 | -34.069 | -9.6163
5.445 | -34.0693 | -9.6271
5.4511 | -34.0696 | -9.6372
5.4572 | -34.0699 | -9.6477
5.4633 | -34.0702 | -9.658
5.4694 | -34.0705 | -9.6684
5.4755 | -34.0708 | -9.6785
5.4816 | -34.0711 | -9.6889
5.4877 | -34.0714 | -9.6992
5.4938 | -34.0717 | -9.7092
5.4999 | -34.072 | -9.7195
5.506 | -34.0723 | -9.7298
5.5121 | -34.0726 | -9.7401
5.5183 | -34.0729 | -9.7503
5.5244 | -34.0732 | -9.7609
5.5305 | -34.0735 | -9.7712
5.5366 | -34.0738 | -9.7813
5.5427 | -34.0741 | -9.7918
5.5488 | -34.0744 | -9.8022
5.5549 | -34.0746 | -9.8126
5.561 | -34.075 | -9.8232
5.5671 | -34.0753 | -9.8334
5.5732 | -34.0756 | -9.8441
5.5793 | -34.0759 | -9.8548
5.5854 | -34.0763 | -9.8651
5.5915 | -34.0765 | -9.8758
5.5976 | -34.0768 | -9.8862
5.6037 | -34.0771 | -9.8967
5.6098 | -34.0774 | -9.907
5.6159 | -34.0777 | -9.9172
5.622 | -34.078 | -9.9277
5.6281 | -34.0783 | -9.9382
5.6342 | -34.0786 | -9.9489
5.6403 | -34.0789 | -9.9593
5.6464 | -34.0792 | -9.9699
5.6525 | -34.0795 | -9.9804
5.6586 | -34.0799 | -9.9905
5.6648 | -34.0802 | -10.0009
5.6709 | -34.0805 | -10.0114
5.677 | -34.0808 | -10.0219
5.6831 | -34.0811 | -10.0324
5.6892 | -34.0814 | -10.043
5.6953 | -34.0817 | -10.0534
5.7014 | -34.082 | -10.0639
5.7075 | -34.0823 | -10.0743
5.7136 | -34.0826 | -10.0846
5.7197 | -34.0829 | -10.095
5.7258 | -34.0832 | -10.1052
5.7319 | -34.0835 | -10.1157
5.738 | -34.0838 | -10.1261
5.7441 | -34.0842 | -10.1365
5.7502 | -34.0845 | -10.1466
5.7563 | -34.0848 | -10.157
5.7624 | -34.0851 | -10.1676
5.7685 | -34.0854 | -10.178
5.7746 | -34.0857 | -10.1887
5.7807 | -34.086 | -10.1992
5.7868 | -34.0863 | -10.2098
5.7929 | -34.0865 | -10.2201
5.799 | -34.0869 | -10.2304
5.8052 | -34.0871 | -10.2406
5.8113 | -34.0874 | -10.2509
5.8174 | -34.0877 | -10.2612
5.8235 | -34.0881 | -10.2715
5.8296 | -34.0883 | -10.2821
5.8357 | -34.0886 | -10.2925
5.8418 | -34.0889 | -10.303
5.8479 | -34.0892 | -10.3135
5.854 | -34.0895 | -10.324
5.8601 | -34.0898 | -10.3343
5.8662 | -34.0901 | -10.3448
5.8723 | -34.0904 | -10.355
5.8784 | -34.0907 | -10.3653
5.8845 | -34.091 | -10.3756
5.8906 | -34.0912 | -10.386
5.8967 | -34.0915 | -10.3962
5.9028 | -34.0919 | -10.4064
5.9089 | -34.0922 | -10.4166
5.915 | -34.0925 | -10.4268
5.9211 | -34.0928 | -10.4371
5.9272 | -34.0931 | -10.4472
5.9333 | -34.0934 | -10.4575
5.9394 | -34.0937 | -10.4678
5.9455 | -34.094 | -10.478
5.9517 | -34.0944 | -10.4881
5.9578 | -34.0947 | -10.4986
5.9639 | -34.095 | -10.5086
5.97 | -34.0954 | -10.519
5.9761 | -34.0957 | -10.5292
5.9822 | -34.096 | -10.5396
5.9883 | -34.0963 | -10.5499
5.9944 | -34.0966 | -10.5603
6.0005 | -34.0969 | -10.5708
6.0066 | -34.0972 | -10.5809
6.0127 | -34.0976 | -10.591
6.0188 | -34.0979 | -10.6016
6.0249 | -34.0982 | -10.612
6.031 | -34.0985 | -10.6221
6.0371 | -34.0988 | -10.6324
6.0432 | -34.0991 | -10.6426
6.0493 | -34.0994 | -10.6527
6.0554 | -34.0998 | -10.6628
6.0615 | -34.1001 | -10.6731
6.0676 | -34.1004 | -10.6836
6.0737 | -34.1007 | -10.694
6.0798 | -34.101 | -10.7043
6.0859 | -34.1014 | -10.7145
6.0921 | -34.1017 | -10.7245
6.0982 | -34.102 | -10.7348
6.1043 | -34.1023 | -10.745
6.1104 | -34.1026 | -10.7554
6.1165 | -34.103 | -10.7652
6.1226 | -34.1033 | -10.7754
6.1287 | -34.1036 | -10.7856
6.1348 | -34.1039 | -10.7958
6.1409 | -34.1042 | -10.806
6.147 | -34.1045 | -10.8162
6.1531 | -34.1048 | -10.8263
6.1592 | -34.1052 | -10.8364
6.1653 | -34.1054 | -10.8465
6.1714 | -34.1057 | -10.8567
6.1775 | -34.1061 | -10.8671
6.1836 | -34.1064 | -10.8773
6.1897 | -34.1067 | -10.8875
6.1958 | -34.107 | -10.898
6.2019 | -34.1074 | -10.9083
6.208 | -34.1077 | -10.9185
6.2141 | -34.108 | -10.9287
6.2202 | -34.1084 | -10.939
6.2263 | -34.1087 | -10.9491
6.2325 | -34.109 | -10.9589
6.2386 | -34.1093 | -10.9688
6.2447 | -34.1096 | -10.9793
6.2508 | -34.11 | -10.9895
6.2569 | -34.1103 | -10.9996
6.263 | -34.1106 | -11.0098
6.2691 | -34.1109 | -11.0198
6.2752 | -34.1112 | -11.0301
6.2813 | -34.1115 | -11.0402
6.2874 | -34.1118 | -11.0503
6.2935 | -34.1121 | -11.0604
6.2996 | -34.1124 | -11.0705
6.3057 | -34.1127 | -11.0807
6.3118 | -34.113 | -11.091
6.3179 | -34.1133 | -11.1011
6.324 | -34.1136 | -11.1112
6.3301 | -34.114 | -11.1213
6.3362 | -34.1143 | -11.1314
6.3423 | -34.1146 | -11.1417
6.3484 | -34.115 | -11.1519
6.3545 | -34.1153 | -11.1617
6.3606 | -34.1156 | -11.1716
6.3667 | -34.1158 | -11.1815
6.3728 | -34.1162 | -11.1917
6.379 | -34.1165 | -11.2018
6.3851 | -34.1168 | -11.2117
6.3912 | -34.1171 | -11.2218
6.3973 | -34.1174 | -11.2319
6.4034 | -34.1177 | -11.2418
6.4095 | -34.118 | -11.2519
6.4156 | -34.1183 | -11.262
6.4217 | -34.1186 | -11.272
6.4278 | -34.1189 | -11.2822
6.4339 | -34.1192 | -11.2922
6.44 | -34.1196 | -11.3022
6.4461 | -34.1199 | -11.3126
6.4522 | -34.1202 | -11.3224
6.4583 | -34.1205 | -11.3326
6.4644 | -34.1208 | -11.3428
6.4705 | -34.1211 | -11.3525
6.4766 | -34.1214 | -11.3626
6.4827 | -34.1217 | -11.3728
6.4888 | -34.122 | -11.3826
6.4949 | -34.1223 | -11.3924
6.501 | -34.1226 | -11.4026
6.5071 | -34.1229 | -11.4124
6.5132 | -34.1232 | -11.4225
6.5194 | -34.1235 | -11.4326
6.5255 | -34.1238 | -11.4424
6.5316 | -34.1241 | -11.4526
6.5377 | -34.1244 | -11.4626
6.5438 | -34.1247 | -11.4724
6.5499 | -34.125 | -11.4823
6.556 | -34.1253 | -11.4923
6.5621 | -34.1256 | -11.5026
6.5682 | -34.1259 | -11.5127
6.5743 | -34.1263 | -11.5226
6.5804 | -34.1266 | -11.5328
6.5865 | -34.1269 | -11.5428
6.5926 | -34.1272 | -11.5527
6.5987 | -34.1275 | -11.5624
6.6048 | -34.1278 | -11.5725
6.6109 | -34.1282 | -11.5824
6.617 | -34.1284 | -11.5926
6.6231 | -34.1287 | -11.6025
6.6292 | -34.1291 | -11.6125
6.6353 | -34.1294 | -11.6228
6.6414 | -34.1297 | -11.6327
6.6475 | -34.13 | -11.6427
6.6536 | -34.1303 | -11.6526
6.6597 | -34.1306 | -11.6623
6.6659 | -34.1309 | -11.6723
6.672 | -34.1312 | -11.6822
6.6781 | -34.1315 | -11.6922
6.6842 | -34.1319 | -11.702
6.6903 | -34.1322 | -11.7119
6.6964 | -34.1325 | -11.722
6.7025 | -34.1328 | -11.7317
6.7086 | -34.1331 | -11.7417
6.7147 | -34.1334 | -11.7517
6.7208 | -34.1337 | -11.7618
6.7269 | -34.134 | -11.7717
6.733 | -34.1343 | -11.7814
6.7391 | -34.1346 | -11.7914
6.7452 | -34.135 | -11.8011
6.7513 | -34.1353 | -11.8107
6.7574 | -34.1356 | -11.8208
6.7635 | -34.1359 | -11.8305
6.7696 | -34.1362 | -11.8403
6.7757 | -34.1365 | -11.8502
6.7818 | -34.1369 | -11.8602
6.7879 | -34.1372 | -11.8699
6.794 | -34.1375 | -11.8796
6.8001 | -34.1378 | -11.8895
6.8063 | -34.1381 | -11.8994
6.8124 | -34.1384 | -11.9094
6.8185 | -34.1387 | -11.9195
6.8246 | -34.139 | -11.9292
6.8307 | -34.1393 | -11.939
6.8368 | -34.1396 | -11.9491
6.8429 | -34.1399 | -11.959
6.849 | -34.1402 | -11.969
6.8551 | -34.1406 | -11.9787
6.8612 | -34.1409 | -11.9885
6.8673 | -34.1412 | -11.9984
6.8734 | -34.1415 | -12.0084
6.8795 | -34.1418 | -12.0183
6.8856 | -34.1421 | -12.0282
6.8917 | -34.1424 | -12.038
6.8978 | -34.1427 | -12.0476
6.9039 | -34.143 | -12.0572
6.91 | -34.1434 | -12.0671
6.9161 | -34.1437 | -12.0767
6.9222 | -34.144 | -12.0865
6.9283 | -34.1444 | -12.0961
6.9344 | -34.1447 | -12.106
6.9405 | -34.145 | -12.1158
6.9466 | -34.1454 | -12.1258
6.9528 | -34.1457 | -12.1359
6.9589 | -34.146 | -12.1458
6.965 | -34.1464 | -12.1558
6.9711 | -34.1467 | -12.1657
6.9772 | -34.1471 | -12.1755
6.9833 | -34.1474 | -12.1853
6.9894 | -34.1477 | -12.1954
6.9955 | -34.148 | -12.205
7.0016 | -34.1483 | -12.2148
7.0077 | -34.1487 | -12.2246
7.0138 | -34.149 | -12.2344
7.0199 | -34.1493 | -12.2441
7.026 | -34.1496 | -12.2539
7.0321 | -34.15 | -12.2635
7.0382 | -34.1503 | -12.273
7.0443 | -34.1506 | -12.283
7.0504 | -34.1509 | -12.2929
7.0565 | -34.1513 | -12.3027
7.0626 | -34.1516 | -12.3123
7.0687 | -34.1519 | -12.3217
7.0748 | -34.1523 | -12.3312
7.0809 | -34.1526 | -12.341
7.087 | -34.1529 | -12.3508
7.0932 | -34.1532 | -12.3606
7.0993 | -34.1536 | -12.3704
7.1054 | -34.1539 | -12.38
7.1115 | -34.1542 | -12.3895
7.1176 | -34.1546 | -12.3992
7.1237 | -34.1549 | -12.4089
7.1298 | -34.1552 | -12.4186
7.1359 | -34.1555 | -12.4283
7.142 | -34.1558 | -12.4378
7.1481 | -34.1562 | -12.4475
7.1542 | -34.1565 | -12.4573
7.1603 | -34.1568 | -12.467
7.1664 | -34.1572 | -12.4767
7.1725 | -34.1575 | -12.4865
7.1786 | -34.1579 | -12.4962
7.1847 | -34.1582 | -12.506
7.1908 | -34.1585 | -12.5159
7.1969 | -34.1588 | -12.5258
7.203 | -34.1592 | -12.5353
7.2091 | -34.1595 | -12.545
7.2152 | -34.1598 | -12.5547
7.2213 | -34.1602 | -12.5643
7.2274 | -34.1605 | -12.5739
7.2335 | -34.1608 | -12.5836
7.2397 | -34.1612 | -12.5933
7.2458 | -34.1615 | -12.6028
7.2519 | -34.1619 | -12.6125
7.258 | -34.1622 | -12.6224
7.2641 | -34.1625 | -12.6321
7.2702 | -34.1628 | -12.642
7.2763 | -34.1632 | -12.6517
7.2824 | -34.1636 | -12.6616
7.2885 | -34.1639 | -12.6713
7.2946 | -34.1642 | -12.681
7.3007 | -34.1645 | -12.6909
7.3068 | -34.1648 | -12.7008
7.3129 | -34.1652 | -12.7107
7.319 | -34.1656 | -12.7204
7.3251 | -34.1659 | -12.7299
7.3312 | -34.1662 | -12.7397
7.3373 | -34.1666 | -12.7492
7.3434 | -34.1669 | -12.7587
7.3495 | -34.1673 | -12.7684
7.3556 | -34.1676 | -12.7782
7.3617 | -34.168 | -12.788
7.3678 | -34.1683 | -12.7977
7.3739 | -34.1686 | -12.8076
7.3801 | -34.1689 | -12.8176
7.3862 | -34.1693 | -12.8273
7.3923 | -34.1697 | -12.8371
7.3984 | -34.17 | -12.8466
7.4045 | -34.1704 | -12.8563
7.4106 | -34.1707 | -12.8659
7.4167 | -34.171 | -12.8757
7.4228 | -34.1714 | -12.8855
7.4289 | -34.1718 | -12.8952
7.435 | -34.1721 | -12.9047
7.4411 | -34.1725 | -12.9146
7.4472 | -34.1728 | -12.9243
7.4533 | -34.1731 | -12.9339
7.4594 | -34.1735 | -12.9438
7.4655 | -34.1738 | -12.9534
7.4716 | -34.1742 | -12.9632
7.4777 | -34.1745 | -12.9729
7.4838 | -34.1749 | -12.9825
7.4899 | -34.1752 | -12.9922
7.496 | -34.1755 | -13.0017
7.5021 | -34.1759 | -13.0114
7.5082 | -34.1762 | -13.0211
7.5143 | -34.1766 | -13.0309
7.5204 | -34.1769 | -13.0407
7.5266 | -34.1773 | -13.0503
7.5327 | -34.1776 | -13.0599
7.5388 | -34.1779 | -13.0699
7.5449 | -34.1783 | -13.0798
7.551 | -34.1787 | -13.0894
7.5571 | -34.179 | -13.0992
7.5632 | -34.1794 | -13.109
7.5693 | -34.1797 | -13.1186
7.5754 | -34.1801 | -13.1282
7.5815 | -34.1804 | -13.138
7.5876 | -34.1808 | -13.1479
7.5937 | -34.1812 | -13.1578
7.5998 | -34.1815 | -13.1674
7.6059 | -34.1819 | -13.1774
7.612 | -34.1822 | -13.1871
7.6181 | -34.1826 | -13.1967
7.6242 | -34.1829 | -13.2067
7.6303 | -34.1833 | -13.2167
7.6364 | -34.1836 | -13.2265
7.6425 | -34.184 | -13.2363
7.6486 | -34.1843 | -13.246
7.6547 | -34.1847 | -13.2556
7.6608 | -34.185 | -13.2654
7.667 | -34.1854 | -13.2752
7.6731 | -34.1857 | -13.2851
7.6792 | -34.1861 | -13.295
7.6853 | -34.1864 | -13.3049
7.6914 | -34.1868 | -13.3148
7.6975 | -34.1872 | -13.3246
7.7036 | -34.1875 | -13.3344
7.7097 | -34.1879 | -13.3443
7.7158 | -34.1882 | -13.3542
7.7219 | -34.1886 | -13.3638
7.728 | -34.189 | -13.3737
7.7341 | -34.1893 | -13.3837
7.7402 | -34.1897 | -13.3935
7.7463 | -34.19 | -13.4033
7.7524 | -34.1904 | -13.4134
7.7585 | -34.1907 | -13.4232
7.7646 | -34.1911 | -13.433
7.7707 | -34.1914 | -13.443
7.7768 | -34.1918 | -13.4529
7.7829 | -34.1921 | -13.463
7.789 | -34.1925 | -13.4731
7.7951 | -34.1928 | -13.4829
7.8012 | -34.1932 | -13.4927
7.8073 | -34.1936 | -13.5028
7.8135 | -34.1939 | -13.5126
7.8196 | -34.1943 | -13.5226
7.8257 | -34.1946 | -13.5327
7.8318 | -34.195 | -13.5426
7.8379 | -34.1954 | -13.5524
7.844 | -34.1957 | -13.5622
7.8501 | -34.1961 | -13.5721
7.8562 | -34.1964 | -13.5822
7.8623 | -34.1968 | -13.592
7.8684 | -34.1971 | -13.6019
7.8745 | -34.1975 | -13.6118
7.8806 | -34.1978 | -13.6219
7.8867 | -34.1982 | -13.6315
7.8928 | -34.1986 | -13.6411
7.8989 | -34.1989 | -13.651
7.905 | -34.1993 | -13.6609
7.9111 | -34.1996 | -13.6709
7.9172 | -34.2 | -13.6808
7.9233 | -34.2003 | -13.6907
7.9294 | -34.2007 | -13.7006
7.9355 | -34.2011 | -13.7103
7.9416 | -34.2014 | -13.7202
7.9477 | -34.2017 | -13.73
7.9539 | -34.2021 | -13.7398
7.96 | -34.2025 | -13.7498
7.9661 | -34.2028 | -13.7597
7.9722 | -34.2032 | -13.7694
7.9783 | -34.2036 | -13.7796
7.9844 | -34.204 | -13.7894
7.9905 | -34.2043 | -13.7995
7.9966 | -34.2047 | -13.8095
8.0027 | -34.205 | -13.8194
8.0088 | -34.2054 | -13.8293
8.0149 | -34.2057 | -13.8392
8.021 | -34.206 | -13.8492
8.0271 | -34.2064 | -13.8594
8.0332 | -34.2068 | -13.8692
8.0393 | -34.2072 | -13.8792
8.0454 | -34.2075 | -13.8894
8.0515 | -34.2079 | -13.8993
8.0576 | -34.2083 | -13.9094
8.0637 | -34.2086 | -13.9193
8.0698 | -34.2089 | -13.929
8.0759 | -34.2093 | -13.9388
8.082 | -34.2096 | -13.9486
8.0881 | -34.21 | -13.9586
8.0942 | -34.2104 | -13.9686
8.1004 | -34.2108 | -13.9786
8.1065 | -34.2111 | -13.9884
8.1126 | -34.2114 | -13.9983
8.1187 | -34.2118 | -14.0083
8.1248 | -34.2122 | -14.018
8.1309 | -34.2125 | -14.0278
8.137 | -34.2129 | -14.0375
8.1431 | -34.2132 | -14.0477
8.1492 | -34.2136 | -14.0575
8.1553 | -34.214 | -14.0675
8.1614 | -34.2143 | -14.0776
8.1675 | -34.2147 | -14.0875
8.1736 | -34.215 | -14.0976
8.1797 | -34.2153 | -14.1076
8.1858 | -34.2157 | -14.1175
8.1919 | -34.2161 | -14.1274
8.198 | -34.2164 | -14.1372
8.2041 | -34.2168 | -14.1471
8.2102 | -34.2172 | -14.157
8.2163 | -34.2175 | -14.1668
8.2224 | -34.2179 | -14.1767
8.2285 | -34.2182 | -14.1863
8.2346 | -34.2186 | -14.1962
8.2408 | -34.219 | -14.2059
8.2469 | -34.2193 | -14.2159
8.253 | -34.2197 | -14.2257
8.2591 | -34.2201 | -14.2356
8.2652 | -34.2205 | -14.2456
8.2713 | -34.2208 | -14.2555
8.2774 | -34.2212 | -14.2655
8.2835 | -34.2216 | -14.2757
8.2896 | -34.2219 | -14.2854
8.2957 | -34.2223 | -14.2954
8.3018 | -34.2227 | -14.3054
8.3079 | -34.223 | -14.3151
8.314 | -34.2234 | -14.325
8.3201 | -34.2238 | -14.3348
8.3262 | -34.2241 | -14.3447
8.3323 | -34.2245 | -14.3546
8.3384 | -34.2249 | -14.3647
8.3445 | -34.2252 | -14.3745
8.3506 | -34.2256 | -14.3845
8.3567 | -34.226 | -14.3945
8.3628 | -34.2263 | -14.4043
8.3689 | -34.2267 | -14.4141
8.375 | -34.2271 | -14.4241
8.3812 | -34.2274 | -14.4339
8.3873 | -34.2278 | -14.4438
8.3934 | -34.2281 | -14.4539
8.3995 | -34.2285 | -14.4636
8.4056 | -34.2289 | -14.4734
8.4117 | -34.2292 | -14.4834
8.4178 | -34.2296 | -14.4933
8.4239 | -34.2299 | -14.503
8.43 | -34.2303 | -14.513
8.4361 | -34.2306 | -14.523
8.4422 | -34.231 | -14.5327
8.4483 | -34.2313 | -14.5425
8.4544 | -34.2317 | -14.5524
8.4605 | -34.2321 | -14.5621
8.4666 | -34.2324 | -14.5723
8.4727 | -34.2328 | -14.5822
8.4788 | -34.2331 | -14.592
8.4849 | -34.2335 | -14.6019
8.491 | -34.2339 | -14.6117
8.4971 | -34.2342 | -14.6213
8.5032 | -34.2346 | -14.6311
8.5093 | -34.235 | -14.6408
8.5154 | -34.2353 | -14.6507
8.5215 | -34.2357 | -14.6605
8.5277 | -34.236 | -14.6704
8.5338 | -34.2364 | -14.6801
8.5399 | -34.2367 | -14.69
8.546 | -34.2372 | -14.6997
8.5521 | -34.2375 | -14.7095
8.5582 | -34.2378 | -14.7193
8.5643 | -34.2382 | -14.7288
8.5704 | -34.2386 | -14.7384
8.5765 | -34.239 | -14.7484
8.5826 | -34.2393 | -14.7578
8.5887 | -34.2397 | -14.7676
8.5948 | -34.2401 | -14.7774
8.6009 | -34.2404 | -14.787
8.607 | -34.2408 | -14.7969
8.6131 | -34.2412 | -14.8068
8.6192 | -34.2416 | -14.8163
8.6253 | -34.2419 | -14.8261
8.6314 | -34.2423 | -14.8358
8.6375 | -34.2426 | -14.8456
8.6436 | -34.243 | -14.8554
8.6497 | -34.2433 | -14.8651
8.6558 | -34.2437 | -14.8748
8.6619 | -34.244 | -14.8846
8.6681 | -34.2444 | -14.8945
8.6742 | -34.2448 | -14.9042
8.6803 | -34.2451 | -14.914
8.6864 | -34.2455 | -14.9234
8.6925 | -34.2459 | -14.9331
8.6986 | -34.2462 | -14.9429
8.7047 | -34.2466 | -14.9528
8.7108 | -34.2469 | -14.9626
8.7169 | -34.2473 | -14.9723
8.723 | -34.2477 | -14.9818
8.7291 | -34.248 | -14.9914
8.7352 | -34.2484 | -15.0013
8.7413 | -34.2487 | -15.0111
8.7474 | -34.2491 | -15.0207
8.7535 | -34.2495 | -15.0305
8.7596 | -34.2499 | -15.0402
8.7657 | -34.2502 | -15.05
8.7718 | -34.2506 | -15.0598
8.7779 | -34.251 | -15.0696
8.784 | -34.2513 | -15.0794
8.7901 | -34.2517 | -15.089
8.7962 | -34.2521 | -15.0985
8.8023 | -34.2524 | -15.1082
8.8084 | -34.2528 | -15.1178
8.8146 | -34.2532 | -15.1275
8.8207 | -34.2535 | -15.1371
8.8268 | -34.2539 | -15.1464
8.8329 | -34.2543 | -15.1561
8.839 | -34.2547 | -15.1658
8.8451 | -34.255 | -15.1754
8.8512 | -34.2554 | -15.185
8.8573 | -34.2557 | -15.1947
8.8634 | -34.2561 | -15.2042
8.8695 | -34.2565 | -15.2141
8.8756 | -34.2568 | -15.2237
8.8817 | -34.2572 | -15.2332
8.8878 | -34.2576 | -15.2431
8.8939 | -34.2579 | -15.2528
8.9 | -34.2583 | -15.2624
8.9061 | -34.2587 | -15.272
8.9122 | -34.259 | -15.2817
8.9183 | -34.2594 | -15.2914
8.9244 | -34.2598 | -15.3012
8.9305 | -34.2602 | -15.3106
8.9366 | -34.2606 | -15.3202
8.9427 | -34.261 | -15.33
8.9488 | -34.2613 | -15.3396
8.955 | -34.2617 | -15.3492
8.9611 | -34.2621 | -15.359
8.9672 | -34.2625 | -15.3685
8.9733 | -34.2629 | -15.378
8.9794 | -34.2633 | -15.3875
8.9855 | -34.2636 | -15.3972
8.9916 | -34.264 | -15.4071
8.9977 | -34.2644 | -15.4167
9.0038 | -34.2648 | -15.4261
9.0099 | -34.2652 | -15.4357
9.016 | -34.2655 | -15.4455
9.0221 | -34.2659 | -15.4549
9.0282 | -34.2663 | -15.4644
9.0343 | -34.2667 | -15.474
9.0404 | -34.2671 | -15.4837
9.0465 | -34.2675 | -15.4934
9.0526 | -34.2679 | -15.5029
9.0587 | -34.2683 | -15.5126
9.0648 | -34.2687 | -15.522
9.0709 | -34.269 | -15.5317
9.077 | -34.2695 | -15.5413
9.0831 | -34.2698 | -15.5511
9.0892 | -34.2702 | -15.5607
9.0953 | -34.2706 | -15.5704
9.1015 | -34.271 | -15.58
9.1076 | -34.2714 | -15.5894
9.1137 | -34.2718 | -15.599
9.1198 | -34.2722 | -15.6085
9.1259 | -34.2726 | -15.6181
9.132 | -34.2729 | -15.6278
9.1381 | -34.2733 | -15.6372
9.1442 | -34.2737 | -15.6465
9.1503 | -34.2741 | -15.6562
9.1564 | -34.2745 | -15.6657
9.1625 | -34.2749 | -15.6753
9.1686 | -34.2753 | -15.6848
9.1747 | -34.2757 | -15.6944
9.1808 | -34.2761 | -15.7038
9.1869 | -34.2765 | -15.7133
9.193 | -34.2769 | -15.7229
9.1991 | -34.2773 | -15.7325
9.2052 | -34.2777 | -15.7421
9.2113 | -34.2781 | -15.7517
9.2174 | -34.2785 | -15.7611
9.2235 | -34.2789 | -15.7707
9.2296 | -34.2793 | -15.7802
9.2357 | -34.2797 | -15.7899
9.2419 | -34.2801 | -15.7994
9.248 | -34.2806 | -15.8089
9.2541 | -34.281 | -15.8185
9.2602 | -34.2814 | -15.8282
9.2663 | -34.2818 | -15.8375
9.2724 | -34.2822 | -15.847
9.2785 | -34.2825 | -15.8565
9.2846 | -34.283 | -15.8658
9.2907 | -34.2834 | -15.8754
9.2968 | -34.2838 | -15.885
9.3029 | -34.2842 | -15.8947
9.309 | -34.2846 | -15.9043
9.3151 | -34.2849 | -15.9137
9.3212 | -34.2854 | -15.9231
9.3273 | -34.2858 | -15.9326
9.3334 | -34.2862 | -15.942
9.3395 | -34.2866 | -15.9513
9.3456 | -34.287 | -15.9607
9.3517 | -34.2873 | -15.9702
9.3578 | -34.2877 | -15.9797
9.3639 | -34.2882 | -15.9891
9.37 | -34.2886 | -15.9987
9.3761 | -34.289 | -16.0083
9.3822 | -34.2894 | -16.0178
9.3884 | -34.2898 | -16.027
9.3945 | -34.2902 | -16.0366
9.4006 | -34.2906 | -16.0461
9.4067 | -34.291 | -16.0557
9.4128 | -34.2914 | -16.0653
9.4189 | -34.2919 | -16.075
9.425 | -34.2923 | -16.0843
9.4311 | -34.2927 | -16.0938
9.4372 | -34.2931 | -16.1033
9.4433 | -34.2935 | -16.1128
9.4494 | -34.2939 | -16.1221
9.4555 | -34.2944 | -16.1316
9.4616 | -34.2948 | -16.1411
9.4677 | -34.2952 | -16.1505
9.4738 | -34.2956 | -16.16
9.4799 | -34.296 | -16.1695
9.486 | -34.2964 | -16.179
9.4921 | -34.2968 | -16.1885
9.4982 | -34.2973 | -16.1978
9.5043 | -34.2977 | -16.2072
9.5104 | -34.2981 | -16.2165
9.5165 | -34.2985 | -16.226
9.5226 | -34.2989 | -16.2356
9.5288 | -34.2993 | -16.2449
9.5349 | -34.2997 | -16.2544
9.541 | -34.3001 | -16.2636
9.5471 | -34.3005 | -16.2731
9.5532 | -34.3009 | -16.2826
9.5593 | -34.3013 | -16.2921
9.5654 | -34.3017 | -16.3014
9.5715 | -34.3022 | -16.3108
9.5776 | -34.3026 | -16.3205
9.5837 | -34.3029 | -16.3299
9.5898 | -34.3033 | -16.3394
9.5959 | -34.3037 | -16.3489
9.602 | -34.3042 | -16.3582
9.6081 | -34.3046 | -16.3676
9.6142 | -34.305 | -16.377
9.6203 | -34.3054 | -16.3864
9.6264 | -34.3058 | -16.3958
9.6325 | -34.3062 | -16.4053
9.6386 | -34.3066 | -16.4146
9.6447 | -34.307 | -16.4241
9.6508 | -34.3074 | -16.4334
9.6569 | -34.3078 | -16.4428
9.663 | -34.3082 | -16.4523
9.6691 | -34.3087 | -16.4617
9.6753 | -34.3091 | -16.4711
9.6814 | -34.3095 | -16.4804
9.6875 | -34.3099 | -16.4896
9.6936 | -34.3103 | -16.4989
9.6997 | -34.3107 | -16.5084
9.7058 | -34.3112 | -16.5178
9.7119 | -34.3116 | -16.5273
9.718 | -34.312 | -16.5369
9.7241 | -34.3124 | -16.5461
9.7302 | -34.3128 | -16.5557
9.7363 | -34.3133 | -16.565
9.7424 | -34.3137 | -16.5742
9.7485 | -34.3141 | -16.5837
9.7546 | -34.3145 | -16.593
9.7607 | -34.3149 | -16.6026
9.7668 | -34.3153 | -16.6121
9.7729 | -34.3157 | -16.6214
9.779 | -34.3162 | -16.6306
9.7851 | -34.3166 | -16.6402
9.7912 | -34.317 | -16.6495
9.7973 | -34.3175 | -16.659
9.8034 | -34.3179 | -16.6683
9.8095 | -34.3183 | -16.6779
9.8157 | -34.3187 | -16.6873
9.8218 | -34.3192 | -16.6968
9.8279 | -34.3196 | -16.7061
9.834 | -34.3199 | -16.7156
9.8401 | -34.3204 | -16.7249
9.8462 | -34.3208 | -16.7341
9.8523 | -34.3212 | -16.7435
9.8584 | -34.3216 | -16.7529
9.8645 | -34.322 | -16.7623
9.8706 | -34.3224 | -16.7718
9.8767 | -34.3228 | -16.7811
9.8828 | -34.3232 | -16.7904
9.8889 | -34.3236 | -16.7996
9.895 | -34.324 | -16.809
9.9011 | -34.3244 | -16.8184
9.9072 | -34.3249 | -16.8277
9.9133 | -34.3253 | -16.8371
9.9194 | -34.3257 | -16.8464
9.9255 | -34.3261 | -16.8557
9.9316 | -34.3265 | -16.865
9.9377 | -34.3269 | -16.8744
9.9438 | -34.3273 | -16.8838
9.9499 | -34.3277 | -16.8934
9.956 | -34.3281 | -16.9028
9.9622 | -34.3286 | -16.9122
9.9683 | -34.329 | -16.9216
9.9744 | -34.3294 | -16.9311
9.9805 | -34.3298 | -16.9405
9.9866 | -34.3302 | -16.9497
9.9927 | -34.3306 | -16.959
9.9988 | -34.331 | -16.9684
10.0049 | -34.3315 | -16.9777
10.011 | -34.3319 | -16.9872
10.0171 | -34.3323 | -16.9964
10.0232 | -34.3327 | -17.0057
10.0293 | -34.3331 | -17.0151
10.0354 | -34.3336 | -17.0244
10.0415 | -34.334 | -17.0338
10.0476 | -34.3344 | -17.0433
10.0537 | -34.3348 | -17.0527
10.0598 | -34.3353 | -17.0621
10.0659 | -34.3356 | -17.0715
10.072 | -34.3361 | -17.0807
10.0781 | -34.3365 | -17.09
10.0842 | -34.3369 | -17.0993
10.0903 | -34.3373 | -17.1086
10.0964 | -34.3377 | -17.1177
10.1026 | -34.3381 | -17.127
10.1087 | -34.3385 | -17.1363
10.1148 | -34.3389 | -17.1458
10.1209 | -34.3394 | -17.155
10.127 | -34.3397 | -17.1642
10.1331 | -34.3402 | -17.1735
10.1392 | -34.3406 | -17.1826
10.1453 | -34.341 | -17.1919
10.1514 | -34.3414 | -17.2013
10.1575 | -34.3418 | -17.2107
10.1636 | -34.3422 | -17.2198
10.1697 | -34.3427 | -17.2292
10.1758 | -34.3431 | -17.2383
10.1819 | -34.3435 | -17.2476
10.188 | -34.3439 | -17.2569
10.1941 | -34.3443 | -17.2662
10.2002 | -34.3447 | -17.2756
10.2063 | -34.3451 | -17.2849
10.2124 | -34.3455 | -17.2943
10.2185 | -34.3459 | -17.3036
10.2246 | -34.3463 | -17.3129
10.2307 | -34.3467 | -17.3222
10.2368 | -34.3472 | -17.3312
10.2429 | -34.3476 | -17.3405
10.2491 | -34.3479 | -17.35
10.2552 | -34.3484 | -17.3594
10.2613 | -34.3488 | -17.3687
10.2674 | -34.3492 | -17.378
10.2735 | -34.3496 | -17.3875
10.2796 | -34.35 | -17.3969
10.2857 | -34.3504 | -17.4063
10.2918 | -34.3509 | -17.4157
10.2979 | -34.3513 | -17.4249
10.304 | -34.3517 | -17.434
10.3101 | -34.3521 | -17.4433
10.3162 | -34.3525 | -17.4528
10.3223 | -34.3529 | -17.4619
10.3284 | -34.3533 | -17.4711
10.3345 | -34.3537 | -17.4804
10.3406 | -34.3541 | -17.4896
10.3467 | -34.3545 | -17.4989
10.3528 | -34.3549 | -17.508
10.3589 | -34.3553 | -17.5173
10.365 | -34.3557 | -17.5267
10.3711 | -34.3561 | -17.5361
10.3772 | -34.3566 | -17.5453
10.3833 | -34.357 | -17.5546
10.3895 | -34.3574 | -17.5638
10.3956 | -34.3578 | -17.5731
10.4017 | -34.3582 | -17.5823
10.4078 | -34.3586 | -17.5916
10.4139 | -34.359 | -17.6009
10.42 | -34.3594 | -17.6104
10.4261 | -34.3598 | -17.6196
10.4322 | -34.3602 | -17.6289
10.4383 | -34.3606 | -17.6381
10.4444 | -34.361 | -17.6473
10.4505 | -34.3614 | -17.6567
10.4566 | -34.3618 | -17.6658
10.4627 | -34.3622 | -17.6749
10.4688 | -34.3627 | -17.6841
10.4749 | -34.3631 | -17.6934
10.481 | -34.3635 | -17.7027
10.4871 | -34.3639 | -17.7119
10.4932 | -34.3643 | -17.7211
10.4993 | -34.3647 | -17.7303
10.5054 | -34.3651 | -17.7397
10.5115 | -34.3656 | -17.749
10.5176 | -34.366 | -17.7581
10.5237 | -34.3664 | -17.7675
10.5298 | -34.3668 | -17.7767
10.536 | -34.3672 | -17.786
10.5421 | -34.3676 | -17.7951
10.5482 | -34.368 | -17.8045
10.5543 | -34.3684 | -17.8137
10.5604 | -34.3688 | -17.8227
10.5665 | -34.3692 | -17.8318
10.5726 | -34.3696 | -17.8411
10.5787 | -34.37 | -17.8504
10.5848 | -34.3704 | -17.8593
10.5909 | -34.3709 | -17.8687
10.597 | -34.3713 | -17.8781
10.6031 | -34.3717 | -17.8875
10.6092 | -34.3721 | -17.8968
10.6153 | -34.3725 | -17.9061
10.6214 | -34.3729 | -17.9152
10.6275 | -34.3733 | -17.9245
10.6336 | -34.3737 | -17.9338
10.6397 | -34.3741 | -17.943
10.6458 | -34.3745 | -17.9521
10.6519 | -34.3749 | -17.9615
10.658 | -34.3753 | -17.9707
10.6641 | -34.3758 | -17.9801
10.6702 | -34.3762 | -17.9894
10.6764 | -34.3766 | -17.9988
10.6825 | -34.377 | -18.0079
10.6886 | -34.3774 | -18.0169
10.6947 | -34.3778 | -18.0261
10.7008 | -34.3782 | -18.0353
10.7069 | -34.3786 | -18.0447
10.713 | -34.379 | -18.054
10.7191 | -34.3794 | -18.0632
10.7252 | -34.3798 | -18.0722
10.7313 | -34.3802 | -18.0815
10.7374 | -34.3806 | -18.0908
10.7435 | -34.381 | -18.1001
10.7496 | -34.3814 | -18.109
10.7557 | -34.3818 | -18.1183
10.7618 | -34.3823 | -18.1273
10.7679 | -34.3827 | -18.1365
10.774 | -34.3831 | -18.1458
10.7801 | -34.3835 | -18.1551
10.7862 | -34.3839 | -18.1642
10.7923 | -34.3843 | -18.1735
10.7984 | -34.3847 | -18.1828
10.8045 | -34.3851 | -18.1918
10.8106 | -34.3855 | -18.2011
10.8168 | -34.3859 | -18.2102
10.8229 | -34.3863 | -18.2195
10.829 | -34.3867 | -18.2286
10.8351 | -34.3871 | -18.2376
10.8412 | -34.3876 | -18.2467
10.8473 | -34.388 | -18.2558
10.8534 | -34.3884 | -18.265
10.8595 | -34.3888 | -18.2743
10.8656 | -34.3892 | -18.2833
10.8717 | -34.3896 | -18.2925
10.8778 | -34.3901 | -18.3016
10.8839 | -34.3905 | -18.3108
10.89 | -34.3909 | -18.32
10.8961 | -34.3913 | -18.3292
10.9022 | -34.3918 | -18.3385
10.9083 | -34.3922 | -18.3475
10.9144 | -34.3925 | -18.3565
10.9205 | -34.3929 | -18.3657
10.9266 | -34.3933 | -18.3748
10.9327 | -34.3937 | -18.3838
10.9388 | -34.3942 | -18.3928
10.9449 | -34.3946 | -18.4019
10.951 | -34.395 | -18.4111
10.9571 | -34.3954 | -18.4203
10.9633 | -34.3958 | -18.4296
10.9694 | -34.3962 | -18.4389
10.9755 | -34.3966 | -18.4481
10.9816 | -34.397 | -18.4574
10.9877 | -34.3975 | -18.4665
10.9938 | -34.3979 | -18.4758
10.9999 | -34.3983 | -18.4848
11.006 | -34.3987 | -18.4939
11.0121 | -34.3991 | -18.503
11.0182 | -34.3995 | -18.5121
11.0243 | -34.4 | -18.5212
11.0304 | -34.4003 | -18.5303
11.0365 | -34.4008 | -18.5394
11.0426 | -34.4012 | -18.5486
11.0487 | -34.4016 | -18.5578
11.0548 | -34.402 | -18.5667
11.0609 | -34.4024 | -18.5759
11.067 | -34.4028 | -18.5851
11.0731 | -34.4032 | -18.5942
11.0792 | -34.4036 | -18.6032
11.0853 | -34.4041 | -18.6124
11.0914 | -34.4045 | -18.6217
11.0975 | -34.4049 | -18.6309
11.1037 | -34.4053 | -18.6399
11.1098 | -34.4057 | -18.6492
11.1159 | -34.4061 | -18.6582
11.122 | -34.4065 | -18.6673
11.1281 | -34.4069 | -18.6765
11.1342 | -34.4073 | -18.6856
11.1403 | -34.4077 | -18.695
11.1464 | -34.4082 | -18.7042
11.1525 | -34.4086 | -18.7133
11.1586 | -34.409 | -18.7227
11.1647 | -34.4094 | -18.7316
11.1708 | -34.4098 | -18.7408
11.1769 | -34.4102 | -18.75
11.183 | -34.4106 | -18.7589
11.1891 | -34.4111 | -18.7681
11.1952 | -34.4115 | -18.7772
11.2013 | -34.4119 | -18.7863
11.2074 | -34.4123 | -18.7954
11.2135 | -34.4127 | -18.8047
11.2196 | -34.4132 | -18.8139
11.2257 | -34.4136 | -18.8233
11.2318 | -34.414 | -18.8322
11.2379 | -34.4145 | -18.8415
11.244 | -34.4149 | -18.8509
11.2502 | -34.4153 | -18.8599
11.2563 | -34.4158 | -18.8692
11.2624 | -34.4162 | -18.8783
11.2685 | -34.4166 | -18.8872
11.2746 | -34.417 | -18.8961
11.2807 | -34.4174 | -18.9054
11.2868 | -34.4179 | -18.9146
11.2929 | -34.4183 | -18.9234
11.299 | -34.4187 | -18.9325
11.3051 | -34.4192 | -18.9418
11.3112 | -34.4196 | -18.9509
11.3173 | -34.42 | -18.9596
11.3234 | -34.4205 | -18.9686
11.3295 | -34.4209 | -18.9777
11.3356 | -34.4213 | -18.9869
11.3417 | -34.4217 | -18.9958
11.3478 | -34.4221 | -19.0051
11.3539 | -34.4224 | -19.0141
11.36 | -34.4229 | -19.0231
11.3661 | -34.4233 | -19.0319
11.3722 | -34.4238 | -19.0408
11.3783 | -34.4242 | -19.05
11.3844 | -34.4247 | -19.0591
11.3906 | -34.4251 | -19.0682
11.3967 | -34.4255 | -19.0773
11.4028 | -34.426 | -19.0864
11.4089 | -34.4264 | -19.0955
11.415 | -34.4268 | -19.1047
11.4211 | -34.4272 | -19.1138
11.4272 | -34.4276 | -19.123
11.4333 | -34.428 | -19.1318
11.4394 | -34.4284 | -19.1409
11.4455 | -34.4288 | -19.15
11.4516 | -34.4293 | -19.1592
11.4577 | -34.4297 | -19.1683
11.4638 | -34.4302 | -19.1773
11.4699 | -34.4306 | -19.1865
11.476 | -34.431 | -19.1958
11.4821 | -34.4315 | -19.2049
11.4882 | -34.4319 | -19.2142
11.4943 | -34.4324 | -19.2234
11.5004 | -34.4328 | -19.2323
11.5065 | -34.4332 | -19.2412
11.5126 | -34.4336 | -19.2502
11.5187 | -34.4341 | -19.2594
11.5248 | -34.4345 | -19.2687
11.5309 | -34.4349 | -19.2779
11.5371 | -34.4354 | -19.2868
11.5432 | -34.4358 | -19.296
11.5493 | -34.4362 | -19.3052
11.5554 | -34.4367 | -19.3143
11.5615 | -34.4371 | -19.3235
11.5676 | -34.4375 | -19.3326
11.5737 | -34.438 | -19.3417
11.5798 | -34.4384 | -19.3506
11.5859 | -34.4389 | -19.3596
11.592 | -34.4393 | -19.3686
11.5981 | -34.4397 | -19.3776
11.6042 | -34.4402 | -19.3867
11.6103 | -34.4406 | -19.3957
11.6164 | -34.441 | -19.4045
11.6225 | -34.4415 | -19.4136
11.6286 | -34.4419 | -19.4226
11.6347 | -34.4423 | -19.4318
11.6408 | -34.4428 | -19.4407
11.6469 | -34.4432 | -19.4497
11.653 | -34.4436 | -19.4586
11.6591 | -34.444 | -19.4677
11.6652 | -34.4445 | -19.4767
11.6713 | -34.4449 | -19.4856
11.6775 | -34.4454 | -19.4946
11.6836 | -34.4458 | -19.5036
11.6897 | -34.4463 | -19.5127
11.6958 | -34.4467 | -19.5219
11.7019 | -34.4471 | -19.531
11.708 | -34.4476 | -19.54
11.7141 | -34.4481 | -19.549
11.7202 | -34.4485 | -19.5581
11.7263 | -34.4489 | -19.5671
11.7324 | -34.4493 | -19.5764
11.7385 | -34.4497 | -19.5855
11.7446 | -34.4501 | -19.5946
11.7507 | -34.4506 | -19.6034
11.7568 | -34.451 | -19.6124
11.7629 | -34.4514 | -19.6216
11.769 | -34.4519 | -19.6306
11.7751 | -34.4523 | -19.6396
11.7812 | -34.4528 | -19.6485
11.7873 | -34.4532 | -19.6576
11.7934 | -34.4536 | -19.6666
11.7995 | -34.4541 | -19.6756
11.8056 | -34.4545 | -19.6846
11.8117 | -34.455 | -19.6936
11.8178 | -34.4554 | -19.7025
11.824 | -34.4559 | -19.7114
11.8301 | -34.4563 | -19.7204
11.8362 | -34.4568 | -19.7293
11.8423 | -34.4572 | -19.7382
11.8484 | -34.4576 | -19.7473
11.8545 | -34.4581 | -19.7561
11.8606 | -34.4586 | -19.765
11.8667 | -34.459 | -19.7742
11.8728 | -34.4594 | -19.783
11.8789 | -34.4599 | -19.7919
11.885 | -34.4603 | -19.801
11.8911 | -34.4608 | -19.81
11.8972 | -34.4612 | -19.8188
11.9033 | -34.4617 | -19.8278
11.9094 | -34.4621 | -19.8367
11.9155 | -34.4625 | -19.8456
11.9216 | -34.463 | -19.8546
11.9277 | -34.4635 | -19.8635
11.9338 | -34.4639 | -19.8725
11.9399 | -34.4643 | -19.8815
11.946 | -34.4648 | -19.8905
11.9521 | -34.4652 | -19.8994
11.9582 | -34.4657 | -19.9083
11.9644 | -34.4661 | -19.9171
11.9705 | -34.4666 | -19.926
11.9766 | -34.4671 | -19.935
11.9827 | -34.4676 | -19.9439
11.9888 | -34.468 | -19.9527
11.9949 | -34.4685 | -19.9615
12.001 | -34.4689 | -19.9705
12.0071 | -34.4694 | -19.9793
12.0132 | -34.4698 | -19.9881
12.0193 | -34.4703 | -19.9971
12.0254 | -34.4707 | -20.006
12.0315 | -34.4712 | -20.0148
12.0376 | -34.4716 | -20.0237
12.0437 | -34.472 | -20.0326
12.0498 | -34.4725 | -20.0416
12.0559 | -34.473 | -20.0504
12.062 | -34.4735 | -20.0591
12.0681 | -34.4739 | -20.0681
12.0742 | -34.4744 | -20.077
12.0803 | -34.4748 | -20.0861
12.0864 | -34.4753 | -20.095
12.0925 | -34.4757 | -20.1039
12.0986 | -34.4762 | -20.1127
12.1047 | -34.4767 | -20.1216
12.1109 | -34.4771 | -20.1303
12.117 | -34.4776 | -20.1391
12.1231 | -34.4781 | -20.148
12.1292 | -34.4785 | -20.1569
12.1353 | -34.479 | -20.1657
12.1414 | -34.4795 | -20.1746
12.1475 | -34.4799 | -20.1836
12.1536 | -34.4803 | -20.1925
12.1597 | -34.4808 | -20.2014
12.1658 | -34.4813 | -20.2102
12.1719 | -34.4817 | -20.2187
12.178 | -34.4822 | -20.2277
12.1841 | -34.4826 | -20.2366
12.1902 | -34.4831 | -20.2455
12.1963 | -34.4836 | -20.2544
12.2024 | -34.4841 | -20.2631
12.2085 | -34.4845 | -20.2719
12.2146 | -34.485 | -20.2809
12.2207 | -34.4855 | -20.2897
12.2268 | -34.4859 | -20.2986
12.2329 | -34.4864 | -20.3074
12.239 | -34.4869 | -20.3159
12.2451 | -34.4874 | -20.3248
12.2513 | -34.4878 | -20.3338
12.2574 | -34.4883 | -20.3426
12.2635 | -34.4887 | -20.3514
12.2696 | -34.4892 | -20.3602
12.2757 | -34.4897 | -20.3687
12.2818 | -34.4901 | -20.3776
12.2879 | -34.4906 | -20.3864
12.294 | -34.4911 | -20.3951
12.3001 | -34.4915 | -20.4039
12.3062 | -34.492 | -20.4127
12.3123 | -34.4925 | -20.4213
12.3184 | -34.493 | -20.4299
12.3245 | -34.4934 | -20.4386
12.3306 | -34.4939 | -20.4473
12.3367 | -34.4944 | -20.456
12.3428 | -34.4948 | -20.4647
12.3489 | -34.4953 | -20.4732
12.355 | -34.4958 | -20.482
12.3611 | -34.4963 | -20.4909
12.3672 | -34.4967 | -20.4996
12.3733 | -34.4972 | -20.5084
12.3794 | -34.4977 | -20.517
12.3855 | -34.4982 | -20.5257
12.3916 | -34.4986 | -20.5345
12.3978 | -34.4991 | -20.5434
12.4039 | -34.4996 | -20.5523
12.41 | -34.5001 | -20.561
12.4161 | -34.5006 | -20.5694
12.4222 | -34.501 | -20.5781
12.4283 | -34.5015 | -20.5868
12.4344 | -34.502 | -20.5957
12.4405 | -34.5024 | -20.6045
12.4466 | -34.5029 | -20.6131
12.4527 | -34.5034 | -20.6218
12.4588 | -34.5039 | -20.6306
12.4649 | -34.5043 | -20.6393
12.471 | -34.5048 | -20.6482
12.4771 | -34.5053 | -20.657
12.4832 | -34.5057 | -20.6658
12.4893 | -34.5062 | -20.6744
12.4954 | -34.5067 | -20.6831
12.5015 | -34.5072 | -20.6919
12.5076 | -34.5077 | -20.7006
12.5137 | -34.5082 | -20.7094
12.5198 | -34.5086 | -20.7179
12.5259 | -34.5091 | -20.7267
12.532 | -34.5096 | -20.7354
12.5382 | -34.5101 | -20.7442
12.5443 | -34.5105 | -20.7529
12.5504 | -34.511 | -20.7617
12.5565 | -34.5115 | -20.7701
12.5626 | -34.512 | -20.779
12.5687 | -34.5125 | -20.7877
12.5748 | -34.5129 | -20.7965
12.5809 | -34.5134 | -20.8051
12.587 | -34.5139 | -20.8139
12.5931 | -34.5144 | -20.8224
12.5992 | -34.5149 | -20.8311
12.6053 | -34.5154 | -20.8398
12.6114 | -34.5159 | -20.8484
12.6175 | -34.5164 | -20.8572
12.6236 | -34.5168 | -20.8659
12.6297 | -34.5173 | -20.8745
12.6358 | -34.5178 | -20.883
12.6419 | -34.5183 | -20.8916
12.648 | -34.5188 | -20.9004
12.6541 | -34.5192 | -20.9091
12.6602 | -34.5198 | -20.9173
12.6663 | -34.5202 | -20.9259
12.6724 | -34.5207 | -20.9345
12.6785 | -34.5212 | -20.9431
12.6847 | -34.5217 | -20.9516
12.6908 | -34.5222 | -20.9602
12.6969 | -34.5227 | -20.9687
12.703 | -34.5231 | -20.9774
12.7091 | -34.5236 | -20.986
12.7152 | -34.5241 | -20.9945
12.7213 | -34.5246 | -21.0032
12.7274 | -34.5251 | -21.0119
12.7335 | -34.5256 | -21.0202
12.7396 | -34.526 | -21.0287
12.7457 | -34.5265 | -21.0375
12.7518 | -34.527 | -21.046
12.7579 | -34.5275 | -21.0548
12.764 | -34.5279 | -21.0635
12.7701 | -34.5284 | -21.0721
12.7762 | -34.5289 | -21.0806
12.7823 | -34.5294 | -21.0892
12.7884 | -34.5299 | -21.0978
12.7945 | -34.5304 | -21.1066
12.8006 | -34.5309 | -21.115
12.8067 | -34.5314 | -21.1236
12.8128 | -34.5319 | -21.132
12.8189 | -34.5324 | -21.1406
12.8251 | -34.5328 | -21.1492
12.8312 | -34.5333 | -21.1579
12.8373 | -34.5338 | -21.1664
12.8434 | -34.5343 | -21.1751
12.8495 | -34.5348 | -21.1836
12.8556 | -34.5352 | -21.1921
12.8617 | -34.5357 | -21.2008
12.8678 | -34.5362 | -21.2095
12.8739 | -34.5366 | -21.2179
12.88 | -34.5371 | -21.2265
12.8861 | -34.5376 | -21.2349
12.8922 | -34.5381 | -21.2437
12.8983 | -34.5385 | -21.2523
12.9044 | -34.539 | -21.2607
12.9105 | -34.5395 | -21.2692
12.9166 | -34.54 | -21.2777
12.9227 | -34.5405 | -21.2862
12.9288 | -34.541 | -21.2948
12.9349 | -34.5415 | -21.3033
12.941 | -34.542 | -21.3118
12.9471 | -34.5425 | -21.3204
12.9532 | -34.543 | -21.329
12.9593 | -34.5435 | -21.3378
12.9654 | -34.544 | -21.3464
12.9716 | -34.5445 | -21.3551
12.9777 | -34.545 | -21.3635
12.9838 | -34.5455 | -21.372
12.9899 | -34.546 | -21.3805
12.996 | -34.5464 | -21.389
13.0021 | -34.5469 | -21.3976
13.0082 | -34.5474 | -21.4061
13.0143 | -34.5479 | -21.4142
13.0204 | -34.5484 | -21.4225
13.0265 | -34.5489 | -21.4311
13.0326 | -34.5494 | -21.4396
13.0387 | -34.5499 | -21.4481
13.0448 | -34.5503 | -21.4567
13.0509 | -34.5508 | -21.465
13.057 | -34.5513 | -21.4735
13.0631 | -34.5518 | -21.482
13.0692 | -34.5524 | -21.4904
13.0753 | -34.5529 | -21.4991
13.0814 | -34.5533 | -21.5077
13.0875 | -34.5538 | -21.5162
13.0936 | -34.5543 | -21.5247
13.0997 | -34.5548 | -21.533
13.1058 | -34.5553 | -21.5414
13.112 | -34.5558 | -21.5498
13.1181 | -34.5563 | -21.558
13.1242 | -34.5567 | -21.5666
13.1303 | -34.5573 | -21.5752
13.1364 | -34.5577 | -21.5836
13.1425 | -34.5582 | -21.592
13.1486 | -34.5587 | -21.6004
13.1547 | -34.5592 | -21.6087
13.1608 | -34.5597 | -21.6171
13.1669 | -34.5602 | -21.6255
13.173 | -34.5606 | -21.634
13.1791 | -34.5611 | -21.6423
13.1852 | -34.5616 | -21.6507
13.1913 | -34.5621 | -21.6589
13.1974 | -34.5625 | -21.6675
13.2035 | -34.563 | -21.6759
13.2096 | -34.5635 | -21.6843
13.2157 | -34.564 | -21.6927
13.2218 | -34.5645 | -21.7011
13.2279 | -34.565 | -21.7093
13.234 | -34.5655 | -21.7176
13.2401 | -34.566 | -21.7261
13.2462 | -34.5665 | -21.7345
13.2524 | -34.567 | -21.743
13.2585 | -34.5675 | -21.7512
13.2646 | -34.5679 | -21.7596
13.2707 | -34.5684 | -21.768
13.2768 | -34.5689 | -21.7766
13.2829 | -34.5694 | -21.7852
13.289 | -34.5699 | -21.7934
13.2951 | -34.5704 | -21.8018
13.3012 | -34.5709 | -21.8102
13.3073 | -34.5714 | -21.8186
13.3134 | -34.5718 | -21.827
13.3195 | -34.5723 | -21.8354
13.3256 | -34.5728 | -21.8439
13.3317 | -34.5732 | -21.8521
13.3378 | -34.5737 | -21.8607
13.3439 | -34.5742 | -21.8692
13.35 | -34.5747 | -21.8778
13.3561 | -34.5752 | -21.886
13.3622 | -34.5756 | -21.8943
13.3683 | -34.5761 | -21.9026
13.3744 | -34.5766 | -21.9109
13.3805 | -34.5771 | -21.9193
13.3866 | -34.5775 | -21.9275
13.3927 | -34.578 | -21.9359
13.3989 | -34.5785 | -21.9442
13.405 | -34.579 | -21.9527
13.4111 | -34.5795 | -21.961
13.4172 | -34.5799 | -21.9694
13.4233 | -34.5804 | -21.9779
13.4294 | -34.5809 | -21.9864
13.4355 | -34.5813 | -21.9945
13.4416 | -34.5818 | -22.0029
13.4477 | -34.5823 | -22.0112
13.4538 | -34.5828 | -22.0195
13.4599 | -34.5833 | -22.028
13.466 | -34.5837 | -22.0364
13.4721 | -34.5842 | -22.0446
13.4782 | -34.5847 | -22.0529
13.4843 | -34.5852 | -22.0613
13.4904 | -34.5857 | -22.0696
13.4965 | -34.5861 | -22.078
13.5026 | -34.5866 | -22.0865
13.5087 | -34.5871 | -22.0947
13.5148 | -34.5875 | -22.1032
13.5209 | -34.588 | -22.1118
13.527 | -34.5885 | -22.1201
13.5331 | -34.589 | -22.1284
13.5393 | -34.5894 | -22.1364
13.5454 | -34.5899 | -22.1447
13.5515 | -34.5904 | -22.1532
13.5576 | -34.5909 | -22.1616
13.5637 | -34.5914 | -22.1701
13.5698 | -34.5918 | -22.1784
13.5759 | -34.5923 | -22.1863
13.582 | -34.5927 | -22.1946
13.5881 | -34.5932 | -22.203
13.5942 | -34.5937 | -22.2113
13.6003 | -34.5942 | -22.2195
13.6064 | -34.5947 | -22.2281
13.6125 | -34.5951 | -22.2366
13.6186 | -34.5956 | -22.2449
13.6247 | -34.5961 | -22.2531
13.6308 | -34.5966 | -22.2616
13.6369 | -34.5971 | -22.27
13.643 | -34.5975 | -22.2784
13.6491 | -34.598 | -22.2867
13.6552 | -34.5985 | -22.2949
13.6613 | -34.599 | -22.3032
13.6674 | -34.5994 | -22.3116
13.6735 | -34.5999 | -22.3199
13.6796 | -34.6004 | -22.3282
13.6858 | -34.6008 | -22.3363
13.6919 | -34.6013 | -22.3446
13.698 | -34.6018 | -22.3529
13.7041 | -34.6022 | -22.3615
13.7102 | -34.6027 | -22.3697
13.7163 | -34.6032 | -22.3781
13.7224 | -34.6037 | -22.3864
13.7285 | -34.6041 | -22.3947
13.7346 | -34.6046 | -22.4031
13.7407 | -34.6051 | -22.4115
13.7468 | -34.6056 | -22.4198
13.7529 | -34.606 | -22.4278
13.759 | -34.6064 | -22.4362
13.7651 | -34.6069 | -22.4447
13.7712 | -34.6074 | -22.453
13.7773 | -34.6078 | -22.4614
13.7834 | -34.6083 | -22.4698
13.7895 | -34.6088 | -22.4781
13.7956 | -34.6092 | -22.4864
13.8017 | -34.6097 | -22.4946
13.8078 | -34.6102 | -22.5029
13.8139 | -34.6107 | -22.511
13.82 | -34.6111 | -22.5191
13.8262 | -34.6116 | -22.5273
13.8323 | -34.6121 | -22.5357
13.8384 | -34.6125 | -22.544
13.8445 | -34.613 | -22.5523
13.8506 | -34.6135 | -22.5606
13.8567 | -34.6139 | -22.5688
13.8628 | -34.6144 | -22.5771
13.8689 | -34.6149 | -22.5854
13.875 | -34.6153 | -22.5937
13.8811 | -34.6158 | -22.602
13.8872 | -34.6162 | -22.6102
13.8933 | -34.6167 | -22.6183
13.8994 | -34.6171 | -22.6266
13.9055 | -34.6176 | -22.635
13.9116 | -34.6181 | -22.6432
13.9177 | -34.6185 | -22.6514
13.9238 | -34.619 | -22.6596
13.9299 | -34.6194 | -22.6676
13.936 | -34.6199 | -22.6758
13.9421 | -34.6203 | -22.6839
13.9482 | -34.6208 | -22.6921
13.9543 | -34.6213 | -22.7004
13.9604 | -34.6217 | -22.7087
13.9665 | -34.6222 | -22.7167
13.9727 | -34.6227 | -22.7252
13.9788 | -34.6231 | -22.7334
13.9849 | -34.6236 | -22.7417
13.991 | -34.6241 | -22.7498
13.9971 | -34.6245 | -22.7576
14.0032 | -34.625 | -22.7659
14.0093 | -34.6254 | -22.774
14.0154 | -34.6259 | -22.782
14.0215 | -34.6263 | -22.7902
14.0276 | -34.6268 | -22.7984
14.0337 | -34.6273 | -22.8067
14.0398 | -34.6277 | -22.815
14.0459 | -34.6282 | -22.8231
14.052 | -34.6287 | -22.8315
14.0581 | -34.6291 | -22.8399
14.0642 | -34.6295 | -22.8481
14.0703 | -34.63 | -22.8564
14.0764 | -34.6305 | -22.8645
14.0825 | -34.6309 | -22.8727
14.0886 | -34.6314 | -22.8808
14.0947 | -34.6319 | -22.889
14.1008 | -34.6323 | -22.8972
14.1069 | -34.6327 | -22.9052
14.1131 | -34.6332 | -22.9135
14.1192 | -34.6337 | -22.9217
14.1253 | -34.6341 | -22.9296
14.1314 | -34.6346 | -22.9378
14.1375 | -34.6351 | -22.9459
14.1436 | -34.6355 | -22.954
14.1497 | -34.636 | -22.9622
14.1558 | -34.6365 | -22.9704
14.1619 | -34.6369 | -22.9785
14.168 | -34.6374 | -22.9867
14.1741 | -34.6379 | -22.9946
14.1802 | -34.6383 | -23.0028
14.1863 | -34.6387 | -23.011
14.1924 | -34.6392 | -23.0192
14.1985 | -34.6397 | -23.0273
14.2046 | -34.6402 | -23.0356
14.2107 | -34.6406 | -23.0435
14.2168 | -34.6411 | -23.0518
14.2229 | -34.6415 | -23.06
14.229 | -34.642 | -23.0682
14.2351 | -34.6424 | -23.0764
14.2412 | -34.6429 | -23.0847
14.2473 | -34.6433 | -23.0925
14.2534 | -34.6438 | -23.1005
14.2596 | -34.6442 | -23.1085
14.2657 | -34.6446 | -23.1167
14.2718 | -34.6451 | -23.1249
14.2779 | -34.6456 | -23.1331
14.284 | -34.646 | -23.1411
14.2901 | -34.6465 | -23.1492
14.2962 | -34.6469 | -23.1574
14.3023 | -34.6474 | -23.1655
14.3084 | -34.6478 | -23.1738
14.3145 | -34.6483 | -23.1819
14.3206 | -34.6487 | -23.1902
14.3267 | -34.6492 | -23.1984
14.3328 | -34.6496 | -23.2065
14.3389 | -34.6501 | -23.2146
14.345 | -34.6505 | -23.2227
14.3511 | -34.651 | -23.2308
14.3572 | -34.6515 | -23.2389
14.3633 | -34.6519 | -23.2473
14.3694 | -34.6524 | -23.2554
14.3755 | -34.6528 | -23.2636
14.3816 | -34.6533 | -23.2718
14.3877 | -34.6538 | -23.2796
14.3938 | -34.6542 | -23.2879
14.4 | -34.6547 | -23.2961
14.4061 | -34.6551 | -23.3041
14.4122 | -34.6556 | -23.3124
14.4183 | -34.656 | -23.3204
14.4244 | -34.6564 | -23.3285
14.4305 | -34.6569 | -23.3367
14.4366 | -34.6574 | -23.3448
14.4427 | -34.6578 | -23.353
14.4488 | -34.6583 | -23.3612
14.4549 | -34.6587 | -23.3692
14.461 | -34.6592 | -23.3772
14.4671 | -34.6596 | -23.3854
14.4732 | -34.6601 | -23.3937
14.4793 | -34.6605 | -23.4019
14.4854 | -34.661 | -23.4099
14.4915 | -34.6614 | -23.4178
14.4976 | -34.6619 | -23.4258
14.5037 | -34.6624 | -23.4337
14.5098 | -34.6628 | -23.4419
14.5159 | -34.6633 | -23.45
14.522 | -34.6637 | -23.4581
14.5281 | -34.6642 | -23.4661
14.5342 | -34.6646 | -23.4743
14.5403 | -34.6651 | -23.4823
14.5465 | -34.6655 | -23.4905
14.5526 | -34.666 | -23.4987
14.5587 | -34.6664 | -23.5069
14.5648 | -34.6669 | -23.5147
14.5709 | -34.6674 | -23.5226
14.577 | -34.6678 | -23.5307
14.5831 | -34.6682 | -23.5388
14.5892 | -34.6687 | -23.5469
14.5953 | -34.6691 | -23.5548
14.6014 | -34.6696 | -23.563
14.6075 | -34.67 | -23.5709
14.6136 | -34.6704 | -23.579
14.6197 | -34.6709 | -23.5872
14.6258 | -34.6713 | -23.5953
14.6319 | -34.6718 | -23.6033
14.638 | -34.6722 | -23.6113
14.6441 | -34.6727 | -23.6194
14.6502 | -34.6731 | -23.6276
14.6563 | -34.6736 | -23.6357
14.6624 | -34.6741 | -23.6438
14.6685 | -34.6745 | -23.6518
14.6746 | -34.675 | -23.66
14.6807 | -34.6754 | -23.6679
14.6869 | -34.6759 | -23.6759
14.693 | -34.6764 | -23.6839
14.6991 | -34.6768 | -23.6919
14.7052 | -34.6773 | -23.6998
14.7113 | -34.6778 | -23.7079
14.7174 | -34.6782 | -23.716
14.7235 | -34.6787 | -23.724
14.7296 | -34.6791 | -23.732
14.7357 | -34.6796 | -23.7398
14.7418 | -34.6801 | -23.7479
14.7479 | -34.6805 | -23.7559
14.754 | -34.681 | -23.764
14.7601 | -34.6815 | -23.7719
14.7662 | -34.6819 | -23.7798
14.7723 | -34.6824 | -23.7876
14.7784 | -34.6828 | -23.7958
14.7845 | -34.6833 | -23.8038
14.7906 | -34.6837 | -23.8119
14.7967 | -34.6841 | -23.82
14.8028 | -34.6846 | -23.8279
14.8089 | -34.6851 | -23.8359
14.815 | -34.6855 | -23.8439
14.8211 | -34.686 | -23.8519
14.8272 | -34.6865 | -23.86
14.8334 | -34.6869 | -23.868
14.8395 | -34.6874 | -23.876
14.8456 | -34.6879 | -23.884
14.8517 | -34.6883 | -23.892
14.8578 | -34.6888 | -23.9001
14.8639 | -34.6892 | -23.9083
14.87 | -34.6897 | -23.9164
14.8761 | -34.6901 | -23.9243
14.8822 | -34.6906 | -23.9324
14.8883 | -34.691 | -23.9405
14.8944 | -34.6915 | -23.9484
14.9005 | -34.6919 | -23.9564
14.9066 | -34.6924 | -23.9645
14.9127 | -34.6929 | -23.9724
14.9188 | -34.6933 | -23.9802
14.9249 | -34.6937 | -23.9883
14.931 | -34.6942 | -23.9963
14.9371 | -34.6947 | -24.0043
14.9432 | -34.6951 | -24.0123
14.9493 | -34.6956 | -24.0202
14.9554 | -34.6961 | -24.0282
14.9615 | -34.6965 | -24.0362
14.9676 | -34.6969 | -24.0443
14.9738 | -34.6974 | -24.0522
14.9799 | -34.6978 | -24.0603
14.986 | -34.6983 | -24.0682
14.9921 | -34.6987 | -24.0761
14.9982 | -34.6992 | -24.084
15.0043 | -34.6996 | -24.0919
15.0104 | -34.7001 | -24.1
15.0165 | -34.7006 | -24.1082
15.0226 | -34.701 | -24.1161
15.0287 | -34.7015 | -24.1242
15.0348 | -34.702 | -24.1322
15.0409 | -34.7024 | -24.1402
15.047 | -34.7029 | -24.1484
15.0531 | -34.7034 | -24.1562
15.0592 | -34.7038 | -24.1643
15.0653 | -34.7043 | -24.1724
15.0714 | -34.7047 | -24.1803
15.0775 | -34.7052 | -24.1882
15.0836 | -34.7057 | -24.196
15.0897 | -34.7062 | -24.2039
15.0958 | -34.7066 | -24.2118
15.1019 | -34.7071 | -24.2199
15.108 | -34.7076 | -24.2277
15.1141 | -34.708 | -24.2357
15.1203 | -34.7085 | -24.2439
15.1264 | -34.7089 | -24.2516
15.1325 | -34.7094 | -24.2596
15.1386 | -34.7099 | -24.2675
15.1447 | -34.7103 | -24.2755
15.1508 | -34.7108 | -24.2835
15.1569 | -34.7112 | -24.2915
15.163 | -34.7117 | -24.2992
15.1691 | -34.7122 | -24.307
15.1752 | -34.7126 | -24.3149
15.1813 | -34.7131 | -24.3229
15.1874 | -34.7136 | -24.3309
15.1935 | -34.7141 | -24.3388
15.1996 | -34.7145 | -24.3466
15.2057 | -34.715 | -24.3546
15.2118 | -34.7154 | -24.3625
15.2179 | -34.7159 | -24.3705
15.224 | -34.7164 | -24.3783
15.2301 | -34.7169 | -24.3861
15.2362 | -34.7174 | -24.394
15.2423 | -34.7178 | -24.402
15.2484 | -34.7183 | -24.4099
15.2545 | -34.7187 | -24.4179
15.2607 | -34.7192 | -24.4259
15.2668 | -34.7197 | -24.4337
15.2729 | -34.7202 | -24.4415
15.279 | -34.7206 | -24.4497
15.2851 | -34.7211 | -24.4578
15.2912 | -34.7216 | -24.4657
15.2973 | -34.7221 | -24.4736
15.3034 | -34.7225 | -24.4814
15.3095 | -34.723 | -24.4894
15.3156 | -34.7235 | -24.4973
15.3217 | -34.7239 | -24.5051
15.3278 | -34.7244 | -24.5132
15.3339 | -34.7249 | -24.5209
15.34 | -34.7254 | -24.5287
15.3461 | -34.7258 | -24.5365
15.3522 | -34.7263 | -24.5443
15.3583 | -34.7267 | -24.5525
15.3644 | -34.7272 | -24.5604
15.3705 | -34.7277 | -24.568
15.3766 | -34.7281 | -24.5761
15.3827 | -34.7286 | -24.5838
15.3888 | -34.7291 | -24.5918
15.3949 | -34.7296 | -24.5998
15.401 | -34.73 | -24.6077
15.4072 | -34.7305 | -24.6155
15.4133 | -34.731 | -24.6234
15.4194 | -34.7315 | -24.6314
15.4255 | -34.7319 | -24.6392
15.4316 | -34.7324 | -24.6471
15.4377 | -34.7329 | -24.6551
15.4438 | -34.7334 | -24.6629
15.4499 | -34.7339 | -24.6708
15.456 | -34.7343 | -24.6787
15.4621 | -34.7348 | -24.6866
15.4682 | -34.7353 | -24.6946
15.4743 | -34.7358 | -24.7023
15.4804 | -34.7362 | -24.7102
15.4865 | -34.7366 | -24.718
15.4926 | -34.7371 | -24.7258
15.4987 | -34.7376 | -24.7336
15.5048 | -34.738 | -24.7416
15.5109 | -34.7385 | -24.7494
15.517 | -34.7391 | -24.7574
15.5231 | -34.7395 | -24.7653
15.5292 | -34.74 | -24.7731
15.5353 | -34.7405 | -24.7812
15.5414 | -34.7409 | -24.7889
15.5476 | -34.7415 | -24.7965
15.5537 | -34.7419 | -24.8044
15.5598 | -34.7424 | -24.8126
15.5659 | -34.7429 | -24.8204
15.572 | -34.7434 | -24.8281
15.5781 | -34.7438 | -24.836
15.5842 | -34.7443 | -24.8438
15.5903 | -34.7448 | -24.8516
15.5964 | -34.7453 | -24.8594
15.6025 | -34.7458 | -24.8671
15.6086 | -34.7463 | -24.8749
15.6147 | -34.7467 | -24.8827
15.6208 | -34.7472 | -24.8906
15.6269 | -34.7477 | -24.8983
15.633 | -34.7482 | -24.9063
15.6391 | -34.7486 | -24.9142
15.6452 | -34.7491 | -24.9221
15.6513 | -34.7496 | -24.9297
15.6574 | -34.7501 | -24.9377
15.6635 | -34.7505 | -24.9455
15.6696 | -34.751 | -24.9535
15.6757 | -34.7515 | -24.9615
15.6818 | -34.752 | -24.9694
15.688 | -34.7524 | -24.9772
15.6941 | -34.7529 | -24.9849
15.7002 | -34.7534 | -24.9927
15.7063 | -34.7539 | -25.0006
15.7124 | -34.7544 | -25.0086
15.7185 | -34.7548 | -25.0163
15.7246 | -34.7553 | -25.024
15.7307 | -34.7559 | -25.0318
15.7368 | -34.7563 | -25.0396
15.7429 | -34.7568 | -25.0473
15.749 | -34.7573 | -25.0553
15.7551 | -34.7578 | -25.0632
15.7612 | -34.7583 | -25.0708
15.7673 | -34.7588 | -25.0785
15.7734 | -34.7593 | -25.0862
15.7795 | -34.7598 | -25.0941
15.7856 | -34.7602 | -25.1018
15.7917 | -34.7608 | -25.1095
15.7978 | -34.7612 | -25.1173
15.8039 | -34.7617 | -25.1253
15.81 | -34.7622 | -25.1331
15.8161 | -34.7627 | -25.1409
15.8222 | -34.7632 | -25.1486
15.8283 | -34.7637 | -25.1563
15.8345 | -34.7642 | -25.1641
15.8406 | -34.7646 | -25.1718
15.8467 | -34.7651 | -25.1796
15.8528 | -34.7656 | -25.1875
15.8589 | -34.7661 | -25.1954
15.865 | -34.7666 | -25.2029
15.8711 | -34.7671 | -25.2109
15.8772 | -34.7676 | -25.2186
15.8833 | -34.768 | -25.2264
15.8894 | -34.7685 | -25.2343
15.8955 | -34.769 | -25.2422
15.9016 | -34.7695 | -25.2498
15.9077 | -34.77 | -25.2575
15.9138 | -34.7704 | -25.2656
15.9199 | -34.7709 | -25.2732
15.926 | -34.7714 | -25.2809
15.9321 | -34.7719 | -25.2886
15.9382 | -34.7723 | -25.2964
15.9443 | -34.7728 | -25.3042
15.9504 | -34.7733 | -25.312
15.9565 | -34.7738 | -25.3198
15.9626 | -34.7743 | -25.3276
15.9687 | -34.7747 | -25.3351
15.9749 | -34.7752 | -25.3427
15.981 | -34.7757 | -25.3504
15.9871 | -34.7762 | -25.3584
15.9932 | -34.7767 | -25.3661
15.9993 | -34.7772 | -25.3738
16.0054 | -34.7777 | -25.3814
16.0115 | -34.7782 | -25.3893
16.0176 | -34.7787 | -25.3974
16.0237 | -34.7792 | -25.4054
16.0298 | -34.7796 | -25.4129
16.0359 | -34.7801 | -25.4208
16.042 | -34.7806 | -25.4284
16.0481 | -34.7811 | -25.4362
16.0542 | -34.7816 | -25.4441
16.0603 | -34.782 | -25.4519
16.0664 | -34.7825 | -25.4599
16.0725 | -34.783 | -25.4677
16.0786 | -34.7835 | -25.4754
16.0847 | -34.784 | -25.4833
16.0908 | -34.7845 | -25.4911
16.0969 | -34.7849 | -25.4989
16.103 | -34.7854 | -25.5069
16.1091 | -34.7859 | -25.5146
16.1152 | -34.7864 | -25.5224
16.1214 | -34.7869 | -25.5303
16.1275 | -34.7874 | -25.5379
16.1336 | -34.7879 | -25.5455
16.1397 | -34.7884 | -25.5532
16.1458 | -34.7889 | -25.5608
16.1519 | -34.7893 | -25.5685
16.158 | -34.7898 | -25.5765
16.1641 | -34.7903 | -25.5844
16.1702 | -34.7908 | -25.592
16.1763 | -34.7913 | -25.5998
16.1824 | -34.7917 | -25.6072
16.1885 | -34.7922 | -25.6148
16.1946 | -34.7927 | -25.6223
16.2007 | -34.7932 | -25.6302
16.2068 | -34.7937 | -25.638
16.2129 | -34.7942 | -25.6455
16.219 | -34.7947 | -25.6527
16.2251 | -34.7952 | -25.6601
16.2312 | -34.7957 | -25.668
16.2373 | -34.7962 | -25.6757
16.2434 | -34.7966 | -25.6834
16.2495 | -34.7971 | -25.6912
16.2556 | -34.7977 | -25.6991
16.2618 | -34.7981 | -25.7066
16.2679 | -34.7986 | -25.7143
16.274 | -34.7991 | -25.7221
16.2801 | -34.7996 | -25.7298
16.2862 | -34.8001 | -25.7373
16.2923 | -34.8006 | -25.7449
16.2984 | -34.801 | -25.7526
16.3045 | -34.8015 | -25.7602
16.3106 | -34.802 | -25.7678
16.3167 | -34.8025 | -25.7755
16.3228 | -34.803 | -25.7828
16.3289 | -34.8035 | -25.7905
16.335 | -34.804 | -25.798
16.3411 | -34.8044 | -25.8057
16.3472 | -34.8049 | -25.8135
16.3533 | -34.8054 | -25.8213
16.3594 | -34.8059 | -25.8287
16.3655 | -34.8064 | -25.8363
16.3716 | -34.8069 | -25.8438
16.3777 | -34.8074 | -25.8514
16.3838 | -34.8078 | -25.8588
16.3899 | -34.8084 | -25.8665
16.396 | -34.8089 | -25.8741
16.4021 | -34.8093 | -25.8817
16.4083 | -34.8098 | -25.8894
16.4144 | -34.8103 | -25.8972
16.4205 | -34.8108 | -25.905
16.4266 | -34.8113 | -25.9126
16.4327 | -34.8118 | -25.9203
16.4388 | -34.8123 | -25.9278
16.4449 | -34.8128 | -25.9353
16.451 | -34.8133 | -25.943
16.4571 | -34.8137 | -25.9508
16.4632 | -34.8142 | -25.9582
16.4693 | -34.8147 | -25.9659
16.4754 | -34.8152 | -25.9738
16.4815 | -34.8157 | -25.9813
16.4876 | -34.8162 | -25.9888
16.4937 | -34.8167 | -25.9965
16.4998 | -34.8172 | -26.0042
16.5059 | -34.8177 | -26.0119
16.512 | -34.8182 | -26.0197
16.5181 | -34.8187 | -26.0276
16.5242 | -34.8191 | -26.0353
16.5303 | -34.8196 | -26.0427
16.5364 | -34.8201 | -26.0503
16.5425 | -34.8206 | -26.0578
16.5487 | -34.821 | -26.0652
16.5548 | -34.8215 | -26.0728
16.5609 | -34.822 | -26.0805
16.567 | -34.8225 | -26.088
16.5731 | -34.823 | -26.0955
16.5792 | -34.8235 | -26.1031
16.5853 | -34.824 | -26.1107
16.5914 | -34.8245 | -26.1184
16.5975 | -34.825 | -26.1262
16.6036 | -34.8255 | -26.1336
16.6097 | -34.826 | -26.1412
16.6158 | -34.8264 | -26.1487
16.6219 | -34.8269 | -26.1563
16.628 | -34.8274 | -26.1639
16.6341 | -34.8278 | -26.1716
16.6402 | -34.8283 | -26.1792
16.6463 | -34.8287 | -26.1868
16.6524 | -34.8292 | -26.1946
16.6585 | -34.8297 | -26.2024
16.6646 | -34.8302 | -26.2098
16.6707 | -34.8307 | -26.217
16.6768 | -34.8312 | -26.2247
16.6829 | -34.8317 | -26.2321
16.689 | -34.8322 | -26.2395
16.6952 | -34.8327 | -26.247
16.7013 | -34.8331 | -26.2544
16.7074 | -34.8337 | -26.2617
16.7135 | -34.8342 | -26.2693
16.7196 | -34.8347 | -26.2769
16.7257 | -34.8352 | -26.2843
16.7318 | -34.8357 | -26.292
16.7379 | -34.8362 | -26.2996
16.744 | -34.8367 | -26.3071
16.7501 | -34.8372 | -26.3147
16.7562 | -34.8377 | -26.3221
16.7623 | -34.8382 | -26.3296
16.7684 | -34.8387 | -26.3373
16.7745 | -34.8391 | -26.3449
16.7806 | -34.8397 | -26.3523
16.7867 | -34.8402 | -26.3598
16.7928 | -34.8406 | -26.3673
16.7989 | -34.8411 | -26.3748
16.805 | -34.8416 | -26.3826
16.8111 | -34.8421 | -26.3899
16.8172 | -34.8425 | -26.3974
16.8233 | -34.8431 | -26.4049
16.8294 | -34.8435 | -26.4127
16.8356 | -34.844 | -26.42
16.8417 | -34.8445 | -26.4278
16.8478 | -34.845 | -26.4352
16.8539 | -34.8455 | -26.4426
16.86 | -34.846 | -26.4503
16.8661 | -34.8464 | -26.4578
16.8722 | -34.8469 | -26.4655
16.8783 | -34.8474 | -26.4733
16.8844 | -34.8478 | -26.4808
16.8905 | -34.8484 | -26.4882
16.8966 | -34.8488 | -26.4958
16.9027 | -34.8493 | -26.5035
16.9088 | -34.8498 | -26.5111
16.9149 | -34.8503 | -26.5184
16.921 | -34.8508 | -26.5258
16.9271 | -34.8513 | -26.5336
16.9332 | -34.8519 | -26.5412
16.9393 | -34.8523 | -26.5489
16.9454 | -34.8528 | -26.5564
16.9515 | -34.8533 | -26.5641
16.9576 | -34.8538 | -26.5718
16.9637 | -34.8542 | -26.5795
16.9698 | -34.8547 | -26.5866
16.9759 | -34.8552 | -26.5943
16.9821 | -34.8557 | -26.602
16.9882 | -34.8561 | -26.6093
16.9943 | -34.8566 | -26.6169
17.0004 | -34.8571 | -26.6242
17.0065 | -34.8576 | -26.6317
17.0126 | -34.8581 | -26.6393
17.0187 | -34.8585 | -26.6469
17.0248 | -34.859 | -26.6543
17.0309 | -34.8594 | -26.6617
17.037 | -34.8599 | -26.6692
17.0431 | -34.8604 | -26.6768
17.0492 | -34.8609 | -26.6844
17.0553 | -34.8614 | -26.692
17.0614 | -34.8618 | -26.6994
17.0675 | -34.8623 | -26.7068
17.0736 | -34.8627 | -26.7145
17.0797 | -34.8632 | -26.7219
17.0858 | -34.8637 | -26.7294
17.0919 | -34.8642 | -26.737
17.098 | -34.8647 | -26.7444
17.1041 | -34.8652 | -26.7518
17.1102 | -34.8657 | -26.7594
17.1163 | -34.8662 | -26.7669
17.1225 | -34.8666 | -26.7744
17.1286 | -34.8671 | -26.7821
17.1347 | -34.8676 | -26.7894
17.1408 | -34.8681 | -26.797
17.1469 | -34.8685 | -26.8043
17.153 | -34.869 | -26.8117
17.1591 | -34.8695 | -26.819
17.1652 | -34.8699 | -26.8266
17.1713 | -34.8704 | -26.834
17.1774 | -34.871 | -26.8415
17.1835 | -34.8715 | -26.8488
17.1896 | -34.8719 | -26.8561
17.1957 | -34.8724 | -26.8636
17.2018 | -34.8729 | -26.8709
17.2079 | -34.8734 | -26.8786
17.214 | -34.8739 | -26.8861
17.2201 | -34.8743 | -26.8938
17.2262 | -34.8748 | -26.9012
17.2323 | -34.8753 | -26.9087
17.2384 | -34.8758 | -26.9162
17.2445 | -34.8763 | -26.9238
17.2506 | -34.8768 | -26.9313
17.2567 | -34.8772 | -26.9389
17.2628 | -34.8777 | -26.9463
17.269 | -34.8781 | -26.9534
17.2751 | -34.8786 | -26.9611
17.2812 | -34.8791 | -26.9688
17.2873 | -34.8796 | -26.9765
17.2934 | -34.8801 | -26.9841
17.2995 | -34.8806 | -26.9918
17.3056 | -34.8811 | -26.999
17.3117 | -34.8816 | -27.0064
17.3178 | -34.882 | -27.014
17.3239 | -34.8825 | -27.0215
17.33 | -34.8829 | -27.029
17.3361 | -34.8834 | -27.0365
17.3422 | -34.8839 | -27.0436
17.3483 | -34.8844 | -27.051
17.3544 | -34.8849 | -27.0584
17.3605 | -34.8853 | -27.0656
17.3666 | -34.8858 | -27.0731
17.3727 | -34.8863 | -27.0806
17.3788 | -34.8867 | -27.0874
17.3849 | -34.8872 | -27.0947
17.391 | -34.8877 | -27.1024
17.3971 | -34.8882 | -27.11
17.4032 | -34.8887 | -27.1173
17.4094 | -34.8892 | -27.1245
17.4155 | -34.8897 | -27.1319
17.4216 | -34.8901 | -27.1395
17.4277 | -34.8906 | -27.147
17.4338 | -34.8911 | -27.1542
17.4399 | -34.8916 | -27.1617
17.446 | -34.8921 | -27.1686
17.4521 | -34.8926 | -27.1758
17.4582 | -34.8931 | -27.1833
17.4643 | -34.8936 | -27.1906
17.4704 | -34.894 | -27.1979
17.4765 | -34.8945 | -27.2053
17.4826 | -34.895 | -27.2126
17.4887 | -34.8955 | -27.22
17.4948 | -34.896 | -27.2274
17.5009 | -34.8965 | -27.2347
17.507 | -34.8969 | -27.2419
17.5131 | -34.8974 | -27.2496
17.5192 | -34.8978 | -27.2566
17.5253 | -34.8983 | -27.264
17.5314 | -34.8988 | -27.2714
17.5375 | -34.8993 | -27.2787
17.5436 | -34.8997 | -27.2863
17.5497 | -34.9002 | -27.2938
17.5559 | -34.9006 | -27.3011
17.562 | -34.9012 | -27.3083
17.5681 | -34.9016 | -27.3156
17.5742 | -34.9021 | -27.3231
17.5803 | -34.9026 | -27.3305
17.5864 | -34.9031 | -27.3377
17.5925 | -34.9036 | -27.3452
17.5986 | -34.904 | -27.3526
17.6047 | -34.9045 | -27.3601
17.6108 | -34.9049 | -27.3671
17.6169 | -34.9054 | -27.3743
17.623 | -34.9059 | -27.3814
17.6291 | -34.9063 | -27.3887
17.6352 | -34.9068 | -27.3963
17.6413 | -34.9073 | -27.4039
17.6474 | -34.9077 | -27.4111
17.6535 | -34.9082 | -27.4186
17.6596 | -34.9086 | -27.4257
17.6657 | -34.9091 | -27.4331
17.6718 | -34.9096 | -27.4407
17.6779 | -34.91 | -27.448
17.684 | -34.9105 | -27.4554
17.6901 | -34.911 | -27.4627
17.6963 | -34.9114 | -27.4701
17.7024 | -34.9119 | -27.4773
17.7085 | -34.9124 | -27.4844
17.7146 | -34.9128 | -27.4918
17.7207 | -34.9133 | -27.4991
17.7268 | -34.9139 | -27.5062
17.7329 | -34.9143 | -27.5136
17.739 | -34.9148 | -27.521
17.7451 | -34.9153 | -27.5282
17.7512 | -34.9157 | -27.5358
17.7573 | -34.9162 | -27.5433
17.7634 | -34.9167 | -27.5505
17.7695 | -34.9172 | -27.5577
17.7756 | -34.9176 | -27.5649
17.7817 | -34.9181 | -27.5721
17.7878 | -34.9186 | -27.5798
17.7939 | -34.9191 | -27.5871
17.8 | -34.9196 | -27.5945
17.8061 | -34.9201 | -27.6019
17.8122 | -34.9205 | -27.6093
17.8183 | -34.921 | -27.6168
17.8244 | -34.9214 | -27.6241
17.8305 | -34.9219 | -27.6314
17.8366 | -34.9224 | -27.6383
17.8428 | -34.9229 | -27.6459
17.8489 | -34.9233 | -27.6533
17.855 | -34.9238 | -27.6604
17.8611 | -34.9242 | -27.6677
17.8672 | -34.9247 | -27.675
17.8733 | -34.9252 | -27.6825
17.8794 | -34.9257 | -27.6896
17.8855 | -34.9261 | -27.6971
17.8916 | -34.9266 | -27.7044
17.8977 | -34.9271 | -27.7116
17.9038 | -34.9276 | -27.7184
17.9099 | -34.9281 | -27.7259
17.916 | -34.9286 | -27.7332
17.9221 | -34.929 | -27.7407
17.9282 | -34.9295 | -27.7481
17.9343 | -34.93 | -27.7552
17.9404 | -34.9304 | -27.7625
17.9465 | -34.9309 | -27.7699
17.9526 | -34.9313 | -27.7774
17.9587 | -34.9318 | -27.7848
17.9648 | -34.9323 | -27.7923
17.9709 | -34.9328 | -27.7999
17.977 | -34.9333 | -27.8071
17.9832 | -34.9337 | -27.8144
17.9893 | -34.9341 | -27.8217
17.9954 | -34.9346 | -27.829
18.0015 | -34.935 | -27.8362
18.0076 | -34.9354 | -27.8437
18.0137 | -34.9359 | -27.8513
18.0198 | -34.9363 | -27.8585
18.0259 | -34.9367 | -27.8656
18.032 | -34.9372 | -27.873
18.0381 | -34.9377 | -27.8802
18.0442 | -34.9381 | -27.8873
18.0503 | -34.9386 | -27.8947
18.0564 | -34.9391 | -27.9019
18.0625 | -34.9395 | -27.9094
18.0686 | -34.94 | -27.9166
18.0747 | -34.9405 | -27.9237
18.0808 | -34.9409 | -27.931
18.0869 | -34.9414 | -27.9384
18.093 | -34.9419 | -27.9456
18.0991 | -34.9424 | -27.9526
18.1052 | -34.9429 | -27.96
18.1113 | -34.9433 | -27.967
18.1174 | -34.9437 | -27.9743
18.1236 | -34.9442 | -27.9818
18.1297 | -34.9447 | -27.9887
18.1358 | -34.9452 | -27.9957
18.1419 | -34.9457 | -28.0029
18.148 | -34.9461 | -28.0099
18.1541 | -34.9465 | -28.0171
18.1602 | -34.947 | -28.0241
18.1663 | -34.9474 | -28.0311
18.1724 | -34.9479 | -28.0385
18.1785 | -34.9484 | -28.0454
18.1846 | -34.9489 | -28.0525
18.1907 | -34.9493 | -28.06
18.1968 | -34.9497 | -28.0673
18.2029 | -34.9502 | -28.0746
18.209 | -34.9507 | -28.0818
18.2151 | -34.9511 | -28.0893
18.2212 | -34.9516 | -28.0962
18.2273 | -34.9521 | -28.1033
18.2334 | -34.9526 | -28.1105
18.2395 | -34.953 | -28.1176
18.2456 | -34.9535 | -28.125
18.2517 | -34.9539 | -28.1324
18.2578 | -34.9544 | -28.1394
18.2639 | -34.9549 | -28.1463
18.2701 | -34.9553 | -28.1533
18.2762 | -34.9558 | -28.1605
18.2823 | -34.9563 | -28.1676
18.2884 | -34.9568 | -28.1751
18.2945 | -34.9572 | -28.1818
18.3006 | -34.9577 | -28.1891
18.3067 | -34.9582 | -28.1963
18.3128 | -34.9587 | -28.2037
18.3189 | -34.9591 | -28.2109
18.325 | -34.9596 | -28.218
18.3311 | -34.96 | -28.2254
18.3372 | -34.9605 | -28.2326
18.3433 | -34.961 | -28.2398
18.3494 | -34.9615 | -28.247
18.3555 | -34.962 | -28.2542
18.3616 | -34.9625 | -28.2613
18.3677 | -34.9629 | -28.2682
18.3738 | -34.9634 | -28.2753
18.3799 | -34.9638 | -28.2827
18.386 | -34.9643 | -28.2902
18.3921 | -34.9647 | -28.2973
18.3982 | -34.9652 | -28.3044
18.4043 | -34.9656 | -28.3114
18.4105 | -34.966 | -28.3188
18.4166 | -34.9665 | -28.3258
18.4227 | -34.967 | -28.3331
18.4288 | -34.9675 | -28.3405
18.4349 | -34.9679 | -28.3474
18.441 | -34.9684 | -28.3546
18.4471 | -34.9688 | -28.362
18.4532 | -34.9693 | -28.3694
18.4593 | -34.9697 | -28.3765
18.4654 | -34.9702 | -28.3836
18.4715 | -34.9707 | -28.3907
18.4776 | -34.9712 | -28.3981
18.4837 | -34.9717 | -28.4054
18.4898 | -34.9721 | -28.4126
18.4959 | -34.9726 | -28.4197
18.502 | -34.9731 | -28.4273
18.5081 | -34.9735 | -28.4344
18.5142 | -34.974 | -28.4414
18.5203 | -34.9744 | -28.4487
18.5264 | -34.9748 | -28.4557
18.5325 | -34.9753 | -28.4631
18.5386 | -34.9758 | -28.4704
18.5447 | -34.9762 | -28.4777
18.5508 | -34.9767 | -28.4848
18.557 | -34.9772 | -28.4922
18.5631 | -34.9776 | -28.4993
18.5692 | -34.9781 | -28.5064
18.5753 | -34.9786 | -28.5135
18.5814 | -34.979 | -28.5206
18.5875 | -34.9794 | -28.5277
18.5936 | -34.9799 | -28.5345
18.5997 | -34.9804 | -28.5417
18.6058 | -34.9808 | -28.5487
18.6119 | -34.9813 | -28.5559
18.618 | -34.9817 | -28.5633
18.6241 | -34.9822 | -28.5705
18.6302 | -34.9826 | -28.5778
18.6363 | -34.9831 | -28.585
18.6424 | -34.9835 | -28.5921
18.6485 | -34.9839 | -28.5992
18.6546 | -34.9844 | -28.6065
18.6607 | -34.9848 | -28.6135
18.6668 | -34.9853 | -28.6207
18.6729 | -34.9858 | -28.6282
18.679 | -34.9862 | -28.6352
18.6851 | -34.9867 | -28.6422
18.6912 | -34.9872 | -28.6493
18.6974 | -34.9876 | -28.6564
18.7035 | -34.9881 | -28.6637
18.7096 | -34.9885 | -28.671
18.7157 | -34.989 | -28.678
18.7218 | -34.9895 | -28.6853
18.7279 | -34.99 | -28.6924
18.734 | -34.9905 | -28.6997
18.7401 | -34.991 | -28.7069
18.7462 | -34.9914 | -28.7134
18.7523 | -34.9919 | -28.7207
18.7584 | -34.9924 | -28.728
18.7645 | -34.9928 | -28.7353
18.7706 | -34.9932 | -28.7426
18.7767 | -34.9937 | -28.7497
18.7828 | -34.9941 | -28.7569
18.7889 | -34.9946 | -28.7639
18.795 | -34.9951 | -28.7713
18.8011 | -34.9955 | -28.7784
18.8072 | -34.9959 | -28.7856
18.8133 | -34.9964 | -28.7929
18.8194 | -34.9969 | -28.7999
18.8255 | -34.9973 | -28.8069
18.8316 | -34.9978 | -28.8141
18.8377 | -34.9982 | -28.821
18.8439 | -34.9987 | -28.8281
18.85 | -34.9992 | -28.8352
18.8561 | -34.9996 | -28.8424
18.8622 | -35.0001 | -28.8495
18.8683 | -35.0005 | -28.8565
18.8744 | -35.001 | -28.8635
18.8805 | -35.0015 | -28.8708
18.8866 | -35.002 | -28.8781
18.8927 | -35.0025 | -28.8849
18.8988 | -35.0029 | -28.8919
18.9049 | -35.0033 | -28.899
18.911 | -35.0038 | -28.9062
18.9171 | -35.0043 | -28.9133
18.9232 | -35.0047 | -28.9203
18.9293 | -35.0052 | -28.9273
18.9354 | -35.0057 | -28.9342
18.9415 | -35.0062 | -28.9413
18.9476 | -35.0066 | -28.9483
18.9537 | -35.0071 | -28.9556
18.9598 | -35.0076 | -28.9621
18.9659 | -35.008 | -28.9691
18.972 | -35.0085 | -28.976
18.9781 | -35.009 | -28.9834
18.9843 | -35.0094 | -28.9906
18.9904 | -35.0099 | -28.998
18.9965 | -35.0104 | -29.0053
19.0026 | -35.0109 | -29.0124
19.0087 | -35.0114 | -29.0195
19.0148 | -35.0119 | -29.0267
19.0209 | -35.0123 | -29.0335
19.027 | -35.0128 | -29.0407
19.0331 | -35.0133 | -29.0476
19.0392 | -35.0138 | -29.0545
19.0453 | -35.0143 | -29.0615
19.0514 | -35.0147 | -29.0684
19.0575 | -35.0152 | -29.0755
19.0636 | -35.0157 | -29.0821
19.0697 | -35.0162 | -29.0893
19.0758 | -35.0166 | -29.0963
19.0819 | -35.0171 | -29.1033
19.088 | -35.0176 | -29.1102
19.0941 | -35.018 | -29.1173
19.1002 | -35.0185 | -29.1238
19.1063 | -35.0189 | -29.1307
19.1124 | -35.0194 | -29.1376
19.1185 | -35.0198 | -29.1448
19.1246 | -35.0203 | -29.1518
19.1308 | -35.0208 | -29.1591
19.1369 | -35.0213 | -29.1658
19.143 | -35.0217 | -29.1729
19.1491 | -35.0222 | -29.1801
19.1552 | -35.0227 | -29.1872
19.1613 | -35.0231 | -29.1941
19.1674 | -35.0235 | -29.2013
19.1735 | -35.024 | -29.2082
19.1796 | -35.0244 | -29.2152
19.1857 | -35.0249 | -29.2222
19.1918 | -35.0254 | -29.2294
19.1979 | -35.0259 | -29.2366
19.204 | -35.0263 | -29.2435
19.2101 | -35.0268 | -29.2501
19.2162 | -35.0273 | -29.2572
19.2223 | -35.0278 | -29.2645
19.2284 | -35.0282 | -29.2715
19.2345 | -35.0287 | -29.2787
19.2406 | -35.0292 | -29.2855
19.2467 | -35.0297 | -29.2924
19.2528 | -35.0302 | -29.2996
19.2589 | -35.0306 | -29.3065
19.265 | -35.031 | -29.3135
19.2712 | -35.0315 | -29.3203
19.2773 | -35.032 | -29.327
19.2834 | -35.0325 | -29.3342
19.2895 | -35.0329 | -29.3409
19.2956 | -35.0334 | -29.3481
19.3017 | -35.0339 | -29.3552
19.3078 | -35.0344 | -29.3622
19.3139 | -35.0349 | -29.3696
19.32 | -35.0354 | -29.3767
19.3261 | -35.0359 | -29.3836
19.3322 | -35.0364 | -29.3904
19.3383 | -35.0368 | -29.397
19.3444 | -35.0374 | -29.404
19.3505 | -35.0379 | -29.4107
19.3566 | -35.0383 | -29.4181
19.3627 | -35.0388 | -29.4255
19.3688 | -35.0392 | -29.4322
19.3749 | -35.0397 | -29.4395
19.381 | -35.0402 | -29.4462
19.3871 | -35.0407 | -29.4531
19.3932 | -35.0411 | -29.4604
19.3993 | -35.0415 | -29.4672
19.4054 | -35.042 | -29.4746
19.4115 | -35.0424 | -29.4817
19.4177 | -35.0429 | -29.4883
19.4238 | -35.0434 | -29.4953
19.4299 | -35.0438 | -29.5024
19.436 | -35.0443 | -29.5094
19.4421 | -35.0447 | -29.5164
19.4482 | -35.0452 | -29.5234
19.4543 | -35.0456 | -29.5299
19.4604 | -35.0461 | -29.5368
19.4665 | -35.0465 | -29.5437
19.4726 | -35.047 | -29.5505
19.4787 | -35.0475 | -29.5575
19.4848 | -35.0479 | -29.5645
19.4909 | -35.0484 | -29.5714
19.497 | -35.0488 | -29.5784
19.5031 | -35.0493 | -29.5852
19.5092 | -35.0497 | -29.5919
19.5153 | -35.0502 | -29.5988
19.5214 | -35.0507 | -29.6052
19.5275 | -35.0511 | -29.6121
19.5336 | -35.0516 | -29.6191
19.5397 | -35.052 | -29.6261
19.5458 | -35.0525 | -29.6331
19.5519 | -35.053 | -29.6401
19.5581 | -35.0534 | -29.6469
19.5642 | -35.0539 | -29.654
19.5703 | -35.0544 | -29.6611
19.5764 | -35.0549 | -29.6679
19.5825 | -35.0554 | -29.6747
19.5886 | -35.0558 | -29.6819
19.5947 | -35.0563 | -29.6887
19.6008 | -35.0568 | -29.6955
19.6069 | -35.0572 | -29.7025
19.613 | -35.0577 | -29.7096
19.6191 | -35.0581 | -29.7164
19.6252 | -35.0586 | -29.7235
19.6313 | -35.0591 | -29.7302
19.6374 | -35.0595 | -29.7373
19.6435 | -35.0599 | -29.7445
19.6496 | -35.0604 | -29.7514
19.6557 | -35.0608 | -29.7583
19.6618 | -35.0613 | -29.7653
19.6679 | -35.0617 | -29.7723
19.674 | -35.0622 | -29.7792
19.6801 | -35.0626 | -29.7865
19.6862 | -35.0631 | -29.7936
19.6923 | -35.0636 | -29.8
19.6984 | -35.0641 | -29.8072
19.7046 | -35.0646 | -29.814
19.7107 | -35.065 | -29.8213
19.7168 | -35.0654 | -29.828
19.7229 | -35.0659 | -29.8352
19.729 | -35.0664 | -29.842
19.7351 | -35.0669 | -29.8483
19.7412 | -35.0674 | -29.8553
19.7473 | -35.0679 | -29.8624
19.7534 | -35.0683 | -29.8696
19.7595 | -35.0688 | -29.8765
19.7656 | -35.0692 | -29.8836
19.7717 | -35.0697 | -29.8901
19.7778 | -35.0701 | -29.8968
19.7839 | -35.0706 | -29.9036
19.79 | -35.0711 | -29.9106
19.7961 | -35.0716 | -29.9176
19.8022 | -35.072 | -29.9242
19.8083 | -35.0725 | -29.9315
19.8144 | -35.073 | -29.9387
19.8205 | -35.0735 | -29.946
19.8266 | -35.074 | -29.9526
19.8327 | -35.0743 | -29.9596
19.8388 | -35.0748 | -29.9661
19.845 | -35.0753 | -29.973
19.8511 | -35.0758 | -29.9799
19.8572 | -35.0763 | -29.9862
19.8633 | -35.0768 | -29.9929
19.8694 | -35.0773 | -29.9997
19.8755 | -35.0779 | -30.0064
19.8816 | -35.0784 | -30.0137
19.8877 | -35.0789 | -30.0208
19.8938 | -35.0794 | -30.0278
19.8999 | -35.0798 | -30.0348
19.906 | -35.0803 | -30.0417
19.9121 | -35.0809 | -30.0481
19.9182 | -35.0814 | -30.0554
19.9243 | -35.0819 | -30.0621
19.9304 | -35.0823 | -30.069
19.9365 | -35.0827 | -30.0759
19.9426 | -35.0832 | -30.0825
19.9487 | -35.0836 | -30.0892
19.9548 | -35.084 | -30.0962
19.9609 | -35.0845 | -30.103
19.967 | -35.085 | -30.1098
19.9731 | -35.0855 | -30.1167
19.9792 | -35.086 | -30.1232
19.9853 | -35.0865 | -30.1298
19.9915 | -35.0869 | -30.1366
19.9976 | -35.0874 | -30.1435
20.0037 | -35.0879 | -30.1504
20.0098 | -35.0884 | -30.1576
20.0159 | -35.0888 | -30.1642
20.022 | -35.0893 | -30.1712
20.0281 | -35.0897 | -30.1779
20.0342 | -35.0902 | -30.1847
20.0403 | -35.0908 | -30.1915
20.0464 | -35.0912 | -30.1984
20.0525 | -35.0917 | -30.205
20.0586 | -35.0922 | -30.2114
20.0647 | -35.0926 | -30.2185
20.0708 | -35.0931 | -30.2258
20.0769 | -35.0936 | -30.2325
20.083 | -35.0941 | -30.2387
20.0891 | -35.0946 | -30.2454
20.0952 | -35.0951 | -30.2522
20.1013 | -35.0956 | -30.2593
20.1074 | -35.096 | -30.2663
20.1135 | -35.0965 | -30.273
20.1196 | -35.097 | -30.2794
20.1257 | -35.0975 | -30.286
20.1319 | -35.098 | -30.2929
20.138 | -35.0984 | -30.2993
20.1441 | -35.0989 | -30.3063
20.1502 | -35.0993 | -30.3128
20.1563 | -35.0997 | -30.3189
20.1624 | -35.1002 | -30.3256
20.1685 | -35.1007 | -30.3323
20.1746 | -35.1012 | -30.3394
20.1807 | -35.1016 | -30.3461
20.1868 | -35.1021 | -30.3529
20.1929 | -35.1026 | -30.3593
20.199 | -35.1031 | -30.3659
20.2051 | -35.1036 | -30.3723
20.2112 | -35.1041 | -30.3789
20.2173 | -35.1045 | -30.3857
20.2234 | -35.105 | -30.3923
20.2295 | -35.1055 | -30.3982
20.2356 | -35.106 | -30.405
20.2417 | -35.1064 | -30.4116
20.2478 | -35.1069 | -30.4186
20.2539 | -35.1074 | -30.4251
20.26 | -35.1079 | -30.4314
20.2661 | -35.1084 | -30.4385
20.2723 | -35.1089 | -30.4451
20.2784 | -35.1093 | -30.4517
20.2845 | -35.1098 | -30.4581
20.2906 | -35.1102 | -30.4646
20.2967 | -35.1107 | -30.4713
20.3028 | -35.1111 | -30.4781
20.3089 | -35.1115 | -30.4849
20.315 | -35.1121 | -30.4914
20.3211 | -35.1125 | -30.4979
20.3272 | -35.113 | -30.5046
20.3333 | -35.1135 | -30.5113
20.3394 | -35.1139 | -30.5178
20.3455 | -35.1144 | -30.5247
20.3516 | -35.1149 | -30.5314
20.3577 | -35.1153 | -30.538
20.3638 | -35.1158 | -30.5446
20.3699 | -35.1163 | -30.5504
20.376 | -35.1168 | -30.5572
20.3821 | -35.1172 | -30.5636
20.3882 | -35.1177 | -30.5704
20.3943 | -35.1181 | -30.5774
20.4004 | -35.1187 | -30.5844
20.4065 | -35.1192 | -30.5911
20.4126 | -35.1197 | -30.5976
20.4188 | -35.1202 | -30.6041
20.4249 | -35.1207 | -30.6109
20.431 | -35.1212 | -30.6172
20.4371 | -35.1217 | -30.6242
20.4432 | -35.1221 | -30.6308
20.4493 | -35.1226 | -30.6376
20.4554 | -35.1231 | -30.6444
20.4615 | -35.1236 | -30.651
20.4676 | -35.1241 | -30.6581
20.4737 | -35.1247 | -30.6642
20.4798 | -35.1252 | -30.6707
20.4859 | -35.1257 | -30.6775
20.492 | -35.1261 | -30.6837
20.4981 | -35.1265 | -30.6902
20.5042 | -35.127 | -30.6962
20.5103 | -35.1275 | -30.7017
20.5164 | -35.128 | -30.7083
20.5225 | -35.1285 | -30.7153
20.5286 | -35.129 | -30.7217
20.5347 | -35.1295 | -30.7284
20.5408 | -35.13 | -30.7347
20.5469 | -35.1305 | -30.7411
20.553 | -35.131 | -30.7477
20.5592 | -35.1315 | -30.7545
20.5653 | -35.1319 | -30.7613
20.5714 | -35.1323 | -30.7679
20.5775 | -35.1328 | -30.7747
20.5836 | -35.1333 | -30.7818
20.5897 | -35.1338 | -30.7878
20.5958 | -35.1342 | -30.7946
20.6019 | -35.1347 | -30.8009
20.608 | -35.1352 | -30.8075
20.6141 | -35.1357 | -30.8142
20.6202 | -35.1362 | -30.8207
20.6263 | -35.1367 | -30.8269
20.6324 | -35.1372 | -30.8335
20.6385 | -35.1376 | -30.8398
20.6446 | -35.1382 | -30.8465
20.6507 | -35.1386 | -30.8525
20.6568 | -35.1391 | -30.8587
20.6629 | -35.1396 | -30.8649
20.669 | -35.14 | -30.8717
20.6751 | -35.1405 | -30.8782
20.6812 | -35.141 | -30.8845
20.6873 | -35.1414 | -30.891
20.6934 | -35.1419 | -30.8973
20.6995 | -35.1424 | -30.9036
20.7057 | -35.1429 | -30.9099
20.7118 | -35.1434 | -30.9162
20.7179 | -35.1439 | -30.9222
20.724 | -35.1443 | -30.929
20.7301 | -35.1448 | -30.9356
20.7362 | -35.1452 | -30.9423
20.7423 | -35.1457 | -30.9488
20.7484 | -35.1462 | -30.9559
20.7545 | -35.1467 | -30.9622
20.7606 | -35.1472 | -30.9685
20.7667 | -35.1476 | -30.9748
20.7728 | -35.1481 | -30.9813
20.7789 | -35.1486 | -30.9874
20.785 | -35.149 | -30.9937
20.7911 | -35.1494 | -31.0004
20.7972 | -35.1498 | -31.0069
20.8033 | -35.1504 | -31.0137
20.8094 | -35.1508 | -31.0198
20.8155 | -35.1513 | -31.0265
20.8216 | -35.1517 | -31.0332
20.8277 | -35.1522 | -31.0399
20.8338 | -35.1527 | -31.0463
20.8399 | -35.1531 | -31.053
20.8461 | -35.1536 | -31.0592
20.8522 | -35.1542 | -31.0652
20.8583 | -35.1547 | -31.0722
20.8644 | -35.1551 | -31.0789
20.8705 | -35.1555 | -31.0852
20.8766 | -35.1559 | -31.0921
20.8827 | -35.1563 | -31.0984
20.8888 | -35.1568 | -31.1049
20.8949 | -35.1573 | -31.1107
20.901 | -35.1578 | -31.1165
20.9071 | -35.1583 | -31.1233
20.9132 | -35.1588 | -31.1306
20.9193 | -35.1592 | -31.1368
20.9254 | -35.1597 | -31.1434
20.9315 | -35.1602 | -31.1502
20.9376 | -35.1606 | -31.1565
20.9437 | -35.161 | -31.1633
20.9498 | -35.1615 | -31.1693
20.9559 | -35.1621 | -31.1757
20.962 | -35.1626 | -31.1823
20.9681 | -35.1631 | -31.1887
20.9742 | -35.1636 | -31.1953
20.9803 | -35.1639 | -31.2022
20.9864 | -35.1644 | -31.2086
20.9926 | -35.1649 | -31.2153
20.9987 | -35.1654 | -31.2204
21.0048 | -35.1659 | -31.2272
21.0109 | -35.1665 | -31.2335
21.017 | -35.167 | -31.2399
21.0231 | -35.1674 | -31.2464
21.0292 | -35.1679 | -31.2524
21.0353 | -35.1685 | -31.2586
21.0414 | -35.169 | -31.2649
21.0475 | -35.1694 | -31.2714
21.0536 | -35.1699 | -31.2782
21.0597 | -35.1704 | -31.2846
21.0658 | -35.1709 | -31.2914
21.0719 | -35.1714 | -31.2976
21.078 | -35.1719 | -31.304
21.0841 | -35.1724 | -31.3099
21.0902 | -35.173 | -31.3158
21.0963 | -35.1736 | -31.3221
21.1024 | -35.174 | -31.3288
21.1085 | -35.1744 | -31.3351
21.1146 | -35.1749 | -31.3418
21.1207 | -35.1754 | -31.3478
21.1268 | -35.1759 | -31.3544
21.133 | -35.1764 | -31.3608
21.1391 | -35.1769 | -31.367
21.1452 | -35.1774 | -31.3728
21.1513 | -35.1779 | -31.3787
21.1574 | -35.1785 | -31.3852
21.1635 | -35.1789 | -31.391
21.1696 | -35.1794 | -31.3972
21.1757 | -35.1799 | -31.4038
21.1818 | -35.1804 | -31.4099
21.1879 | -35.1808 | -31.4168
21.194 | -35.1813 | -31.4234
21.2001 | -35.1817 | -31.4291
21.2062 | -35.1823 | -31.4357
21.2123 | -35.1827 | -31.4416
21.2184 | -35.1833 | -31.448
21.2245 | -35.1838 | -31.4542
21.2306 | -35.1843 | -31.4602
21.2367 | -35.1848 | -31.4667
21.2428 | -35.1853 | -31.4733
21.2489 | -35.1857 | -31.4791
21.255 | -35.1862 | -31.4852
21.2611 | -35.1867 | -31.491
21.2672 | -35.1871 | -31.4975
21.2733 | -35.1875 | -31.5038
21.2795 | -35.188 | -31.5102
21.2856 | -35.1886 | -31.5168
21.2917 | -35.189 | -31.523
21.2978 | -35.1895 | -31.5288
21.3039 | -35.19 | -31.5344
21.31 | -35.1903 | -31.541
21.3161 | -35.1908 | -31.5464
21.3222 | -35.1912 | -31.5533
21.3283 | -35.1916 | -31.5594
21.3344 | -35.192 | -31.5655
21.3405 | -35.1924 | -31.5715
21.3466 | -35.1929 | -31.5772
21.3527 | -35.1935 | -31.5842
21.3588 | -35.194 | -31.5902
21.3649 | -35.1945 | -31.5963
21.371 | -35.195 | -31.6024
21.3771 | -35.1955 | -31.6079
21.3832 | -35.1959 | -31.6142
21.3893 | -35.1963 | -31.6199
21.3954 | -35.1968 | -31.626
21.4015 | -35.1973 | -31.6324
21.4076 | -35.1977 | -31.639
21.4137 | -35.1982 | -31.6457
21.4199 | -35.1986 | -31.6524
21.426 | -35.199 | -31.6589
21.4321 | -35.1995 | -31.6645
21.4382 | -35.1999 | -31.6714
21.4443 | -35.2003 | -31.6776
21.4504 | -35.2008 | -31.6842
21.4565 | -35.2012 | -31.6892
21.4626 | -35.2017 | -31.6954
21.4687 | -35.2021 | -31.7019
21.4748 | -35.2026 | -31.7073
21.4809 | -35.203 | -31.7135
21.487 | -35.2035 | -31.7197
21.4931 | -35.2039 | -31.7252
21.4992 | -35.2044 | -31.7314
21.5053 | -35.2048 | -31.7376
21.5114 | -35.2053 | -31.7436
21.5175 | -35.2058 | -31.7498
21.5236 | -35.2062 | -31.7558
21.5297 | -35.2067 | -31.7613
21.5358 | -35.2071 | -31.7674
21.5419 | -35.2075 | -31.774
21.548 | -35.208 | -31.7799
21.5541 | -35.2085 | -31.7862
21.5602 | -35.209 | -31.7928
21.5664 | -35.2096 | -31.7986
21.5725 | -35.2101 | -31.8046
21.5786 | -35.2107 | -31.8106
21.5847 | -35.211 | -31.8163
21.5908 | -35.2114 | -31.8224
21.5969 | -35.2119 | -31.8281
21.603 | -35.2124 | -31.8342
21.6091 | -35.2129 | -31.8405
21.6152 | -35.2134 | -31.8468
21.6213 | -35.2138 | -31.853
21.6274 | -35.2142 | -31.8591
21.6335 | -35.2146 | -31.8648
21.6396 | -35.215 | -31.8714
21.6457 | -35.2154 | -31.8775
21.6518 | -35.2158 | -31.8838
21.6579 | -35.2163 | -31.8905
21.664 | -35.2167 | -31.8969
21.6701 | -35.2171 | -31.9022
21.6762 | -35.2176 | -31.9088
21.6823 | -35.2181 | -31.9156
21.6884 | -35.2186 | -31.9221
21.6945 | -35.2192 | -31.9284
21.7006 | -35.2198 | -31.9348
21.7068 | -35.2201 | -31.9408
21.7129 | -35.2206 | -31.9473
21.719 | -35.2211 | -31.9536
21.7251 | -35.2215 | -31.9596
21.7312 | -35.222 | -31.9662
21.7373 | -35.2225 | -31.9726
21.7434 | -35.223 | -31.9786
21.7495 | -35.2234 | -31.9851
21.7556 | -35.2239 | -31.9914
21.7617 | -35.2243 | -31.9974
21.7678 | -35.2248 | -32.0037
21.7739 | -35.2251 | -32.01
21.78 | -35.2256 | -32.0159
21.7861 | -35.2261 | -32.0223
21.7922 | -35.2265 | -32.0281
21.7983 | -35.2269 | -32.0343
21.8044 | -35.2274 | -32.0405
21.8105 | -35.2279 | -32.047
21.8166 | -35.2283 | -32.0533
21.8227 | -35.2288 | -32.0591
21.8288 | -35.2293 | -32.065
21.8349 | -35.2297 | -32.0715
21.841 | -35.2302 | -32.0777
21.8471 | -35.2306 | -32.0837
21.8533 | -35.2311 | -32.0899
21.8594 | -35.2315 | -32.0963
21.8655 | -35.232 | -32.1021
21.8716 | -35.2326 | -32.1078
21.8777 | -35.233 | -32.1137
21.8838 | -35.2335 | -32.1188
21.8899 | -35.234 | -32.1246
21.896 | -35.2345 | -32.1312
21.9021 | -35.2349 | -32.1368
21.9082 | -35.2354 | -32.1428
21.9143 | -35.2359 | -32.1487
21.9204 | -35.2364 | -32.1549
21.9265 | -35.2369 | -32.1609
21.9326 | -35.2374 | -32.1672
21.9387 | -35.2378 | -32.1729
21.9448 | -35.2382 | -32.1793
21.9509 | -35.2388 | -32.1859
21.957 | -35.2393 | -32.1917
21.9631 | -35.2397 | -32.1978
21.9692 | -35.2403 | -32.2038
21.9753 | -35.2408 | -32.2096
21.9814 | -35.2413 | -32.2156
21.9875 | -35.2416 | -32.2216
21.9937 | -35.2421 | -32.2284
21.9998 | -35.2425 | -32.2339
22.0059 | -35.243 | -32.2399
22.012 | -35.2435 | -32.2455
22.0181 | -35.244 | -32.2512
22.0242 | -35.2445 | -32.2571
22.0303 | -35.2449 | -32.2635
22.0364 | -35.2454 | -32.2694
22.0425 | -35.2459 | -32.2759
22.0486 | -35.2463 | -32.2822
22.0547 | -35.2469 | -32.2877
22.0608 | -35.2474 | -32.2935
22.0669 | -35.2478 | -32.2995
22.073 | -35.2482 | -32.3055
22.0791 | -35.2488 | -32.3123
22.0852 | -35.2492 | -32.3183
22.0913 | -35.2496 | -32.3234
22.0974 | -35.25 | -32.3296
22.1035 | -35.2504 | -32.3354
22.1096 | -35.2508 | -32.3413
22.1157 | -35.2512 | -32.3474
22.1218 | -35.2518 | -32.3534
22.1279 | -35.2523 | -32.3586
22.134 | -35.2527 | -32.3647
22.1402 | -35.2532 | -32.3704
22.1463 | -35.2536 | -32.3766
22.1524 | -35.2539 | -32.3825
22.1585 | -35.2544 | -32.3889
22.1646 | -35.2549 | -32.3953
22.1707 | -35.2553 | -32.4015
22.1768 | -35.2558 | -32.4073
22.1829 | -35.2563 | -32.413
22.189 | -35.2567 | -32.4196
22.1951 | -35.2572 | -32.425
22.2012 | -35.2576 | -32.431
22.2073 | -35.258 | -32.4369
22.2134 | -35.2585 | -32.4428
22.2195 | -35.259 | -32.4488
22.2256 | -35.2595 | -32.4548
22.2317 | -35.26 | -32.4605
22.2378 | -35.2605 | -32.467
22.2439 | -35.2609 | -32.4729
22.25 | -35.2614 | -32.4787
22.2561 | -35.2618 | -32.4851
22.2622 | -35.2622 | -32.4909
22.2683 | -35.2626 | -32.4971
22.2744 | -35.2631 | -32.5035
22.2806 | -35.2634 | -32.5093
22.2867 | -35.2638 | -32.5147
22.2928 | -35.2642 | -32.5199
22.2989 | -35.2647 | -32.5263
22.305 | -35.2653 | -32.5314
22.3111 | -35.2658 | -32.5371
22.3172 | -35.2663 | -32.544
22.3233 | -35.2668 | -32.5499
22.3294 | -35.2672 | -32.5552
22.3355 | -35.2676 | -32.5619
22.3416 | -35.2681 | -32.5681
22.3477 | -35.2685 | -32.5741
22.3538 | -35.269 | -32.5799
22.3599 | -35.2694 | -32.586
22.366 | -35.2699 | -32.5912
22.3721 | -35.2703 | -32.5971
22.3782 | -35.2708 | -32.6031
22.3843 | -35.2712 | -32.6085
22.3904 | -35.2716 | -32.6139
22.3965 | -35.272 | -32.62
22.4026 | -35.2724 | -32.6263
22.4087 | -35.2729 | -32.6324
22.4148 | -35.2732 | -32.638
22.4209 | -35.2737 | -32.6438
22.4271 | -35.2743 | -32.6501
22.4332 | -35.2748 | -32.6562
22.4393 | -35.2752 | -32.6624
22.4454 | -35.2756 | -32.6676
22.4515 | -35.276 | -32.6734
22.4576 | -35.2765 | -32.6795
22.4637 | -35.277 | -32.6856
22.4698 | -35.2775 | -32.6914
22.4759 | -35.2779 | -32.6968
22.482 | -35.2783 | -32.7022
22.4881 | -35.2788 | -32.7082
22.4942 | -35.2793 | -32.7146
22.5003 | -35.2797 | -32.7204
22.5064 | -35.2802 | -32.7266
22.5125 | -35.2806 | -32.7324
22.5186 | -35.281 | -32.7381
22.5247 | -35.2814 | -32.7435
22.5308 | -35.2818 | -32.7495
22.5369 | -35.2823 | -32.7551
22.543 | -35.2827 | -32.7609
22.5491 | -35.2832 | -32.7658
22.5552 | -35.2836 | -32.7722
22.5613 | -35.2841 | -32.7778
22.5675 | -35.2846 | -32.7829
22.5736 | -35.2851 | -32.789
22.5797 | -35.2855 | -32.7954
22.5858 | -35.2859 | -32.8019
22.5919 | -35.2864 | -32.8076
22.598 | -35.2868 | -32.814
22.6041 | -35.2872 | -32.8195
22.6102 | -35.2877 | -32.8254
22.6163 | -35.288 | -32.8312
22.6224 | -35.2884 | -32.8369
22.6285 | -35.2889 | -32.8429
22.6346 | -35.2894 | -32.8487
22.6407 | -35.2899 | -32.8546
22.6468 | -35.2903 | -32.8603
22.6529 | -35.2908 | -32.8664
22.659 | -35.2912 | -32.8722
22.6651 | -35.2916 | -32.8773
22.6712 | -35.2921 | -32.8827
22.6773 | -35.2925 | -32.889
22.6834 | -35.2929 | -32.8952
22.6895 | -35.2934 | -32.9009
22.6956 | -35.2938 | -32.907
22.7017 | -35.2943 | -32.9126
22.7079 | -35.2947 | -32.918
22.714 | -35.2951 | -32.924
22.7201 | -35.2956 | -32.9298
22.7262 | -35.2961 | -32.9361
22.7323 | -35.2964 | -32.9423
22.7384 | -35.2969 | -32.948
22.7445 | -35.2974 | -32.9539
22.7506 | -35.2978 | -32.9592
22.7567 | -35.2982 | -32.9648
22.7628 | -35.2986 | -32.9714
22.7689 | -35.299 | -32.9775
22.775 | -35.2994 | -32.9836
22.7811 | -35.2998 | -32.9892
22.7872 | -35.3002 | -32.9951
22.7933 | -35.3007 | -33.001
22.7994 | -35.3011 | -33.0068
22.8055 | -35.3015 | -33.0129
22.8116 | -35.3019 | -33.019
22.8177 | -35.3024 | -33.0248
22.8238 | -35.3028 | -33.0304
22.8299 | -35.3033 | -33.0363
22.836 | -35.3038 | -33.0416
22.8421 | -35.3042 | -33.0474
22.8482 | -35.3046 | -33.0534
22.8544 | -35.305 | -33.0595
22.8605 | -35.3055 | -33.0659
22.8666 | -35.306 | -33.0718
22.8727 | -35.3063 | -33.0774
22.8788 | -35.3068 | -33.0829
22.8849 | -35.3073 | -33.0887
22.891 | -35.3078 | -33.0938
22.8971 | -35.3083 | -33.0995
22.9032 | -35.3087 | -33.1048
22.9093 | -35.3091 | -33.1099
22.9154 | -35.3095 | -33.115
22.9215 | -35.31 | -33.1214
22.9276 | -35.3103 | -33.1268
22.9337 | -35.3108 | -33.1319
22.9398 | -35.3112 | -33.1374
22.9459 | -35.3117 | -33.1436
22.952 | -35.3121 | -33.1497
22.9581 | -35.3125 | -33.1554
22.9642 | -35.3129 | -33.1617
22.9703 | -35.3132 | -33.1676
22.9764 | -35.3137 | -33.1736
22.9825 | -35.3141 | -33.1794
22.9886 | -35.3146 | -33.1861
22.9948 | -35.315 | -33.1924
23.0009 | -35.3153 | -33.198
23.007 | -35.3157 | -33.2045
23.0131 | -35.3161 | -33.2101
23.0192 | -35.3165 | -33.2159
23.0253 | -35.3169 | -33.222
23.0314 | -35.3173 | -33.2276
23.0375 | -35.3177 | -33.233
23.0436 | -35.3181 | -33.2389
23.0497 | -35.3186 | -33.245
23.0558 | -35.319 | -33.2511
23.0619 | -35.3194 | -33.256
23.068 | -35.3199 | -33.2623
23.0741 | -35.3204 | -33.2681
23.0802 | -35.3209 | -33.2741
23.0863 | -35.3213 | -33.2792
23.0924 | -35.3217 | -33.285
23.0985 | -35.322 | -33.291
23.1046 | -35.3225 | -33.2974
23.1107 | -35.3229 | -33.3032
23.1168 | -35.3233 | -33.3091
23.1229 | -35.3237 | -33.3143
23.129 | -35.3242 | -33.3199
23.1351 | -35.3247 | -33.3258
23.1413 | -35.3251 | -33.3318
23.1474 | -35.3256 | -33.3376
23.1535 | -35.326 | -33.3432
23.1596 | -35.3264 | -33.349
23.1657 | -35.3268 | -33.3547
23.1718 | -35.3272 | -33.3605
23.1779 | -35.3275 | -33.3664
23.184 | -35.328 | -33.3716
23.1901 | -35.3284 | -33.3774
23.1962 | -35.3289 | -33.3831
23.2023 | -35.3293 | -33.3893
23.2084 | -35.3297 | -33.3954
23.2145 | -35.33 | -33.4009
23.2206 | -35.3305 | -33.407
23.2267 | -35.3309 | -33.4133
23.2328 | -35.3313 | -33.4192
23.2389 | -35.3317 | -33.4252
23.245 | -35.3322 | -33.4311
23.2511 | -35.3326 | -33.4369
23.2572 | -35.3329 | -33.4429
23.2633 | -35.3334 | -33.4495
23.2694 | -35.3338 | -33.4556
23.2755 | -35.3341 | -33.4614
23.2817 | -35.3345 | -33.4673
23.2878 | -35.335 | -33.4734
23.2939 | -35.3354 | -33.4796
23.3 | -35.3358 | -33.4859
23.3061 | -35.3362 | -33.4916
23.3122 | -35.3366 | -33.497
23.3183 | -35.3371 | -33.5026
23.3244 | -35.3375 | -33.5088
23.3305 | -35.338 | -33.5145
23.3366 | -35.3383 | -33.5206
23.3427 | -35.3388 | -33.5262
23.3488 | -35.3391 | -33.5326
23.3549 | -35.3394 | -33.5386
23.361 | -35.3397 | -33.5445
23.3671 | -35.3402 | -33.5504
23.3732 | -35.3405 | -33.5556
23.3793 | -35.341 | -33.5615
23.3854 | -35.3415 | -33.5678
23.3915 | -35.342 | -33.5732
23.3976 | -35.3424 | -33.5788
23.4037 | -35.3428 | -33.5853
23.4098 | -35.3432 | -33.5917
23.4159 | -35.3437 | -33.5974
23.422 | -35.3441 | -33.6031
23.4282 | -35.3446 | -33.6091
23.4343 | -35.345 | -33.6151
23.4404 | -35.3454 | -33.6209
23.4465 | -35.3457 | -33.6265
23.4526 | -35.3461 | -33.6327
23.4587 | -35.3466 | -33.6378
23.4648 | -35.3469 | -33.6441
23.4709 | -35.3472 | -33.6503
23.477 | -35.3477 | -33.6562
23.4831 | -35.3482 | -33.6625
23.4892 | -35.3485 | -33.6685
23.4953 | -35.3488 | -33.6741
23.5014 | -35.3492 | -33.68
23.5075 | -35.3495 | -33.6859
23.5136 | -35.35 | -33.6921
23.5197 | -35.3504 | -33.6979
23.5258 | -35.3508 | -33.7035
23.5319 | -35.3512 | -33.7094
23.538 | -35.3515 | -33.7151
23.5441 | -35.3519 | -33.7211
23.5502 | -35.3524 | -33.7273
23.5563 | -35.3528 | -33.7333
23.5624 | -35.3533 | -33.739
23.5686 | -35.3536 | -33.7453
23.5747 | -35.354 | -33.751
23.5808 | -35.3544 | -33.7566
23.5869 | -35.3549 | -33.7624
23.593 | -35.3553 | -33.7683
23.5991 | -35.3557 | -33.7744
23.6052 | -35.356 | -33.7791
23.6113 | -35.3564 | -33.7857
23.6174 | -35.3569 | -33.7912
23.6235 | -35.3573 | -33.7969
23.6296 | -35.3576 | -33.8028
23.6357 | -35.3581 | -33.8083
23.6418 | -35.3584 | -33.8146
23.6479 | -35.3589 | -33.8217
23.654 | -35.3593 | -33.8277
23.6601 | -35.3597 | -33.8331
23.6662 | -35.36 | -33.8386
23.6723 | -35.3604 | -33.8448
23.6784 | -35.3607 | -33.8511
23.6845 | -35.3611 | -33.8565
23.6906 | -35.3615 | -33.8622
23.6967 | -35.362 | -33.8673
23.7028 | -35.3625 | -33.874
23.7089 | -35.3628 | -33.8799
23.7151 | -35.3632 | -33.8859
23.7212 | -35.3636 | -33.8921
23.7273 | -35.364 | -33.8989
23.7334 | -35.3644 | -33.9047
23.7395 | -35.3648 | -33.9103
23.7456 | -35.3653 | -33.9167
23.7517 | -35.3656 | -33.9225
23.7578 | -35.366 | -33.9282
23.7639 | -35.3664 | -33.9344
23.77 | -35.3668 | -33.9399
23.7761 | -35.3673 | -33.9456
23.7822 | -35.3675 | -33.952
23.7883 | -35.3679 | -33.9583
23.7944 | -35.3683 | -33.9641
23.8005 | -35.3686 | -33.9702
23.8066 | -35.369 | -33.9764
23.8127 | -35.3695 | -33.983
23.8188 | -35.3699 | -33.9883
23.8249 | -35.3704 | -33.9939
23.831 | -35.3707 | -33.9998
23.8371 | -35.3711 | -34.005
23.8432 | -35.3715 | -34.011
23.8493 | -35.3719 | -34.0179
23.8555 | -35.3722 | -34.024
23.8616 | -35.3727 | -34.0297
23.8677 | -35.3731 | -34.0352
23.8738 | -35.3734 | -34.0414
23.8799 | -35.3738 | -34.0477
23.886 | -35.374 | -34.0533
23.8921 | -35.3744 | -34.0585
23.8982 | -35.3749 | -34.0646
23.9043 | -35.3754 | -34.0705
23.9104 | -35.3758 | -34.0755
23.9165 | -35.3763 | -34.0815
23.9226 | -35.3766 | -34.0875
23.9287 | -35.3771 | -34.0935
23.9348 | -35.3775 | -34.0993
23.9409 | -35.3778 | -34.105
23.947 | -35.3783 | -34.1112
23.9531 | -35.3787 | -34.1169
23.9592 | -35.3791 | -34.123
23.9653 | -35.3796 | -34.1289
23.9714 | -35.3801 | -34.1349
23.9775 | -35.3805 | -34.1409
23.9836 | -35.3809 | -34.1467
23.9897 | -35.3812 | -34.1523
23.9958 | -35.3816 | -34.158
24.002 | -35.3819 | -34.1642
24.0081 | -35.3823 | -34.1704
24.0142 | -35.3827 | -34.1763
24.0203 | -35.3831 | -34.1825
24.0264 | -35.3835 | -34.1885
24.0325 | -35.3839 | -34.1942
24.0386 | -35.3843 | -34.2001
24.0447 | -35.3846 | -34.2059
24.0508 | -35.3849 | -34.2116
24.0569 | -35.3853 | -34.2172
24.063 | -35.3857 | -34.2228
24.0691 | -35.3861 | -34.2282
24.0752 | -35.3865 | -34.2343
24.0813 | -35.3869 | -34.2405
24.0874 | -35.3872 | -34.2463
24.0935 | -35.3877 | -34.2519
24.0996 | -35.388 | -34.2577
24.1057 | -35.3884 | -34.2633
24.1118 | -35.3888 | -34.2701
24.1179 | -35.3892 | -34.2764
24.124 | -35.3895 | -34.2826
24.1301 | -35.3898 | -34.2881
24.1362 | -35.3901 | -34.294
24.1424 | -35.3906 | -34.2999
24.1485 | -35.3911 | -34.3053
24.1546 | -35.3914 | -34.3118
24.1607 | -35.3917 | -34.3178
24.1668 | -35.3922 | -34.3239
24.1729 | -35.3925 | -34.33
24.179 | -35.3928 | -34.3359
24.1851 | -35.3932 | -34.3414
24.1912 | -35.3936 | -34.3468
24.1973 | -35.3939 | -34.3528
24.2034 | -35.3943 | -34.3591
24.2095 | -35.3948 | -34.3659
24.2156 | -35.3951 | -34.3716
24.2217 | -35.3955 | -34.3776
24.2278 | -35.3959 | -34.3836
24.2339 | -35.3962 | -34.3893
24.24 | -35.3966 | -34.3959
24.2461 | -35.3969 | -34.4022
24.2522 | -35.3972 | -34.4077
24.2583 | -35.3974 | -34.4132
24.2644 | -35.3979 | -34.419
24.2705 | -35.3982 | -34.4248
24.2766 | -35.3986 | -34.4314
24.2827 | -35.3991 | -34.4367
24.2889 | -35.3995 | -34.4432
24.295 | -35.3999 | -34.4493
24.3011 | -35.4003 | -34.4547
24.3072 | -35.4006 | -34.4603
24.3133 | -35.401 | -34.4655
24.3194 | -35.4015 | -34.4714
24.3255 | -35.4019 | -34.4772
24.3316 | -35.4023 | -34.483
24.3377 | -35.4027 | -34.4885
24.3438 | -35.403 | -34.4943
24.3499 | -35.4034 | -34.5003
24.356 | -35.4038 | -34.5066
24.3621 | -35.4042 | -34.5129
24.3682 | -35.4046 | -34.5191
24.3743 | -35.4049 | -34.5253
24.3804 | -35.4053 | -34.531
24.3865 | -35.4058 | -34.5375
24.3926 | -35.4063 | -34.5429
24.3987 | -35.4067 | -34.5487
24.4048 | -35.407 | -34.5543
24.4109 | -35.4073 | -34.5605
24.417 | -35.4077 | -34.5667
24.4231 | -35.4081 | -34.5729
24.4293 | -35.4084 | -34.579
24.4354 | -35.4089 | -34.5847
24.4415 | -35.4092 | -34.591
24.4476 | -35.4097 | -34.5968
24.4537 | -35.41 | -34.6022
24.4598 | -35.4104 | -34.6081
24.4659 | -35.4109 | -34.6142
24.472 | -35.4113 | -34.6203
24.4781 | -35.4116 | -34.6261
24.4842 | -35.4121 | -34.6318
24.4903 | -35.4125 | -34.6379
24.4964 | -35.4128 | -34.6432
24.5025 | -35.4132 | -34.6491
24.5086 | -35.4136 | -34.6555
24.5147 | -35.4139 | -34.6613
24.5208 | -35.4144 | -34.6675
24.5269 | -35.4147 | -34.673
24.533 | -35.4152 | -34.6787
24.5391 | -35.4155 | -34.6845
24.5452 | -35.4159 | -34.6902
24.5513 | -35.4162 | -34.6957
24.5574 | -35.4166 | -34.7015
24.5635 | -35.4169 | -34.7076
24.5696 | -35.4172 | -34.7132
24.5758 | -35.4176 | -34.7185
24.5819 | -35.4179 | -34.7245
24.588 | -35.4183 | -34.7304
24.5941 | -35.4187 | -34.7364
24.6002 | -35.4192 | -34.7426
24.6063 | -35.4196 | -34.7484
24.6124 | -35.42 | -34.7537
24.6185 | -35.4204 | -34.7594
24.6246 | -35.4208 | -34.7662
24.6307 | -35.4212 | -34.7722
24.6368 | -35.4215 | -34.7783
24.6429 | -35.4219 | -34.7841
24.649 | -35.4222 | -34.7896
24.6551 | -35.4226 | -34.7955
24.6612 | -35.4229 | -34.8019
24.6673 | -35.4232 | -34.8085
24.6734 | -35.4235 | -34.8146
24.6795 | -35.4239 | -34.821
24.6856 | -35.4243 | -34.8266
24.6917 | -35.4247 | -34.8326
24.6978 | -35.4251 | -34.8385
24.7039 | -35.4255 | -34.8446
24.71 | -35.4258 | -34.8503
24.7162 | -35.4262 | -34.8563
24.7223 | -35.4265 | -34.8619
24.7284 | -35.427 | -34.8681
24.7345 | -35.4274 | -34.8739
24.7406 | -35.4278 | -34.88
24.7467 | -35.4282 | -34.8854
24.7528 | -35.4286 | -34.8911
24.7589 | -35.429 | -34.8969
24.765 | -35.4294 | -34.9031
24.7711 | -35.4298 | -34.9086
24.7772 | -35.4301 | -34.9144
24.7833 | -35.4306 | -34.9202
24.7894 | -35.4309 | -34.9259
24.7955 | -35.4313 | -34.9317
24.8016 | -35.4317 | -34.938
24.8077 | -35.432 | -34.9441
24.8138 | -35.4324 | -34.9503
24.8199 | -35.4328 | -34.9564
24.826 | -35.4332 | -34.962
24.8321 | -35.4336 | -34.9683
24.8382 | -35.434 | -34.9736
24.8443 | -35.4342 | -34.9793
24.8504 | -35.4345 | -34.9848
24.8565 | -35.4349 | -34.9908
24.8627 | -35.4353 | -34.9968
24.8688 | -35.4357 | -35.0034
24.8749 | -35.436 | -35.0091
24.881 | -35.4363 | -35.0149
24.8871 | -35.4368 | -35.0209
24.8932 | -35.4372 | -35.0268
24.8993 | -35.4376 | -35.0331
24.9054 | -35.4381 | -35.0386
24.9115 | -35.4385 | -35.0444
24.9176 | -35.4389 | -35.0504
24.9237 | -35.4392 | -35.0561
24.9298 | -35.4396 | -35.0619
24.9359 | -35.44 | -35.068
24.942 | -35.4404 | -35.0739
24.9481 | -35.4408 | -35.0795
24.9542 | -35.4411 | -35.0854
24.9603 | -35.4415 | -35.0918
24.9664 | -35.4419 | -35.0975
24.9725 | -35.4423 | -35.1033
24.9786 | -35.4427 | -35.1093
24.9847 | -35.443 | -35.1155
24.9908 | -35.4434 | -35.1212
24.9969 | -35.4438 | -35.1266
25.0031 | -35.4442 | -35.1326
25.0092 | -35.4446 | -35.1383
25.0153 | -35.445 | -35.1439
25.0214 | -35.4453 | -35.1497
25.0275 | -35.4457 | -35.156
25.0336 | -35.4461 | -35.1617
25.0397 | -35.4465 | -35.1679
25.0458 | -35.4469 | -35.1737
25.0519 | -35.4473 | -35.1796
25.058 | -35.4476 | -35.1854
25.0641 | -35.4479 | -35.1911
25.0702 | -35.4483 | -35.1973
25.0763 | -35.4487 | -35.2033
25.0824 | -35.4491 | -35.2098
25.0885 | -35.4494 | -35.2158
25.0946 | -35.4498 | -35.2216
25.1007 | -35.4501 | -35.2276
25.1068 | -35.4505 | -35.2331
25.1129 | -35.4508 | -35.2389
25.119 | -35.4512 | -35.2447
25.1251 | -35.4516 | -35.2503
25.1312 | -35.452 | -35.2562
25.1373 | -35.4524 | -35.2623
25.1435 | -35.4528 | -35.2681
25.1496 | -35.4531 | -35.2736
25.1557 | -35.4534 | -35.2795
25.1618 | -35.4538 | -35.2859
25.1679 | -35.4542 | -35.2914
25.174 | -35.4546 | -35.2967
25.1801 | -35.455 | -35.303
25.1862 | -35.4553 | -35.3092
25.1923 | -35.4556 | -35.3156
25.1984 | -35.456 | -35.3215
25.2045 | -35.4563 | -35.3276
25.2106 | -35.4567 | -35.3342
25.2167 | -35.4571 | -35.3396
25.2228 | -35.4574 | -35.3456
25.2289 | -35.4577 | -35.3516
25.235 | -35.4581 | -35.357
25.2411 | -35.4585 | -35.363
25.2472 | -35.4589 | -35.3688
25.2533 | -35.4592 | -35.3749
25.2594 | -35.4596 | -35.3803
25.2655 | -35.4599 | -35.3865
25.2716 | -35.4602 | -35.3916
25.2777 | -35.4606 | -35.3972
25.2838 | -35.461 | -35.4028
25.29 | -35.4614 | -35.4084
25.2961 | -35.4618 | -35.4141
25.3022 | -35.4622 | -35.4199
25.3083 | -35.4626 | -35.4255
25.3144 | -35.4629 | -35.4315
25.3205 | -35.4633 | -35.4372
25.3266 | -35.4637 | -35.4427
25.3327 | -35.4642 | -35.4485
25.3388 | -35.4645 | -35.4547
25.3449 | -35.4649 | -35.4607
25.351 | -35.4653 | -35.4664
25.3571 | -35.4656 | -35.4724
25.3632 | -35.4659 | -35.4785
25.3693 | -35.4663 | -35.4843
25.3754 | -35.4668 | -35.4902
25.3815 | -35.4671 | -35.4956
25.3876 | -35.4676 | -35.5014
25.3937 | -35.468 | -35.5072
25.3998 | -35.4683 | -35.5129
25.4059 | -35.4687 | -35.5192
25.412 | -35.4691 | -35.5251
25.4181 | -35.4694 | -35.5303
25.4242 | -35.4698 | -35.5363
25.4304 | -35.4701 | -35.5423
25.4365 | -35.4705 | -35.5479
25.4426 | -35.4708 | -35.5537
25.4487 | -35.4712 | -35.5594
25.4548 | -35.4715 | -35.5651
25.4609 | -35.4719 | -35.5707
25.467 | -35.4724 | -35.5755
25.4731 | -35.4728 | -35.5816
25.4792 | -35.4731 | -35.587
25.4853 | -35.4735 | -35.5928
25.4914 | -35.4739 | -35.5983
25.4975 | -35.4743 | -35.6042
25.5036 | -35.4746 | -35.6101
25.5097 | -35.4751 | -35.6162
25.5158 | -35.4755 | -35.6217
25.5219 | -35.4758 | -35.6278
25.528 | -35.4761 | -35.6333
25.5341 | -35.4765 | -35.6391
25.5402 | -35.4768 | -35.6448
25.5463 | -35.4772 | -35.6513
25.5524 | -35.4776 | -35.6571
25.5585 | -35.4779 | -35.6633
25.5646 | -35.4784 | -35.6695
25.5707 | -35.4788 | -35.6753
25.5769 | -35.4792 | -35.6814
25.583 | -35.4796 | -35.6872
25.5891 | -35.48 | -35.6929
25.5952 | -35.4805 | -35.6985
25.6013 | -35.4809 | -35.7047
25.6074 | -35.4813 | -35.7102
25.6135 | -35.4816 | -35.7162
25.6196 | -35.482 | -35.7225
25.6257 | -35.4824 | -35.728
25.6318 | -35.4828 | -35.7337
25.6379 | -35.4832 | -35.7394
25.644 | -35.4836 | -35.7454
25.6501 | -35.4839 | -35.7509
25.6562 | -35.4844 | -35.7567
25.6623 | -35.4847 | -35.7624
25.6684 | -35.485 | -35.7683
25.6745 | -35.4854 | -35.7741
25.6806 | -35.4858 | -35.78
25.6867 | -35.4861 | -35.7858
25.6928 | -35.4865 | -35.7918
25.6989 | -35.4869 | -35.7976
25.705 | -35.4872 | -35.8035
25.7111 | -35.4876 | -35.8101
25.7173 | -35.488 | -35.8157
25.7234 | -35.4884 | -35.8219
25.7295 | -35.4888 | -35.8278
25.7356 | -35.4892 | -35.8337
25.7417 | -35.4895 | -35.8391
25.7478 | -35.4899 | -35.8446
25.7539 | -35.4903 | -35.8503
25.76 | -35.4906 | -35.8563
25.7661 | -35.491 | -35.8624
25.7722 | -35.4914 | -35.8683
25.7783 | -35.4918 | -35.8742
25.7844 | -35.4922 | -35.8797
25.7905 | -35.4925 | -35.8852
25.7966 | -35.4929 | -35.8912
25.8027 | -35.4933 | -35.8973
25.8088 | -35.4937 | -35.9039
25.8149 | -35.4941 | -35.9097
25.821 | -35.4944 | -35.9151
25.8271 | -35.4948 | -35.921
25.8332 | -35.4951 | -35.9269
25.8393 | -35.4955 | -35.9327
25.8454 | -35.4958 | -35.9384
25.8515 | -35.4961 | -35.9444
25.8576 | -35.4966 | -35.9501
25.8638 | -35.4969 | -35.9555
25.8699 | -35.4973 | -35.9612
25.876 | -35.4977 | -35.9672
25.8821 | -35.4979 | -35.9733
25.8882 | -35.4984 | -35.9788
25.8943 | -35.4987 | -35.9847
25.9004 | -35.4991 | -35.9903
25.9065 | -35.4995 | -35.9959
25.9126 | -35.4999 | -36.002
25.9187 | -35.5003 | -36.0073
25.9248 | -35.5006 | -36.0129
25.9309 | -35.501 | -36.0184
25.937 | -35.5013 | -36.0241
25.9431 | -35.5016 | -36.0302
25.9492 | -35.502 | -36.0362
25.9553 | -35.5024 | -36.0423
25.9614 | -35.5028 | -36.0478
25.9675 | -35.5031 | -36.0536
25.9736 | -35.5034 | -36.0598
25.9797 | -35.5038 | -36.0654
25.9858 | -35.5042 | -36.0712
25.9919 | -35.5046 | -36.0771
25.998 | -35.505 | -36.0829
26.0042 | -35.5053 | -36.0886
26.0103 | -35.5057 | -36.0943
26.0164 | -35.5061 | -36.1002
26.0225 | -35.5065 | -36.1057
26.0286 | -35.5069 | -36.1114
26.0347 | -35.5073 | -36.1172
26.0408 | -35.5077 | -36.1229
26.0469 | -35.508 | -36.1284
26.053 | -35.5083 | -36.1343
26.0591 | -35.5087 | -36.1403
26.0652 | -35.5091 | -36.146
26.0713 | -35.5093 | -36.1517
26.0774 | -35.5096 | -36.1574
26.0835 | -35.5099 | -36.1635
26.0896 | -35.5103 | -36.1692
26.0957 | -35.5107 | -36.1752
26.1018 | -35.511 | -36.1809
26.1079 | -35.5114 | -36.1869
26.114 | -35.5118 | -36.1925
26.1201 | -35.5121 | -36.1981
26.1262 | -35.5125 | -36.2038
26.1323 | -35.5128 | -36.2094
26.1384 | -35.5133 | -36.2151
26.1445 | -35.5136 | -36.2211
26.1507 | -35.5141 | -36.2269
26.1568 | -35.5144 | -36.2324
26.1629 | -35.5148 | -36.2382
26.169 | -35.5152 | -36.244
26.1751 | -35.5155 | -36.2495
26.1812 | -35.5159 | -36.2549
26.1873 | -35.5162 | -36.2606
26.1934 | -35.5167 | -36.2663
26.1995 | -35.517 | -36.2721
26.2056 | -35.5174 | -36.2779
26.2117 | -35.5178 | -36.2834
26.2178 | -35.5182 | -36.2888
26.2239 | -35.5185 | -36.2946
26.23 | -35.5189 | -36.3004
26.2361 | -35.5193 | -36.3061
26.2422 | -35.5196 | -36.3115
26.2483 | -35.52 | -36.3175
26.2544 | -35.5204 | -36.323
26.2605 | -35.5208 | -36.3284
26.2666 | -35.5211 | -36.3343
26.2727 | -35.5214 | -36.34
26.2788 | -35.5218 | -36.346
26.2849 | -35.5222 | -36.3515
26.2911 | -35.5227 | -36.357
26.2972 | -35.5231 | -36.3631
26.3033 | -35.5235 | -36.3685
26.3094 | -35.5239 | -36.3743
26.3155 | -35.5243 | -36.3795
26.3216 | -35.5247 | -36.3855
26.3277 | -35.525 | -36.3907
26.3338 | -35.5254 | -36.3963
26.3399 | -35.5258 | -36.4018
26.346 | -35.5262 | -36.4069
26.3521 | -35.5265 | -36.413
26.3582 | -35.5269 | -36.4189
26.3643 | -35.5272 | -36.4243
26.3704 | -35.5276 | -36.4301
26.3765 | -35.528 | -36.4357
26.3826 | -35.5283 | -36.4407
26.3887 | -35.5288 | -36.4464
26.3948 | -35.5292 | -36.452
26.4009 | -35.5296 | -36.4573
26.407 | -35.53 | -36.4631
26.4131 | -35.5304 | -36.469
26.4192 | -35.5308 | -36.4743
26.4253 | -35.5311 | -36.4801
26.4314 | -35.5314 | -36.4853
26.4376 | -35.5318 | -36.491
26.4437 | -35.5322 | -36.4971
26.4498 | -35.5325 | -36.5023
26.4559 | -35.533 | -36.5077
26.462 | -35.5334 | -36.5135
26.4681 | -35.5338 | -36.5193
26.4742 | -35.5341 | -36.525
26.4803 | -35.5345 | -36.5306
26.4864 | -35.5349 | -36.5358
26.4925 | -35.5353 | -36.5413
26.4986 | -35.5357 | -36.5472
26.5047 | -35.5361 | -36.5528
26.5108 | -35.5364 | -36.5587
26.5169 | -35.5368 | -36.5641
26.523 | -35.5372 | -36.5694
26.5291 | -35.5376 | -36.5748
26.5352 | -35.5379 | -36.5804
26.5413 | -35.5383 | -36.586
26.5474 | -35.5387 | -36.592
26.5535 | -35.539 | -36.5977
26.5596 | -35.5394 | -36.6026
26.5657 | -35.5398 | -36.6082
26.5718 | -35.5401 | -36.6138
26.578 | -35.5405 | -36.6193
26.5841 | -35.5409 | -36.6248
26.5902 | -35.5412 | -36.6305
26.5963 | -35.5417 | -36.6358
26.6024 | -35.542 | -36.6415
26.6085 | -35.5423 | -36.647
26.6146 | -35.5427 | -36.6526
26.6207 | -35.5431 | -36.6585
26.6268 | -35.5435 | -36.6642
26.6329 | -35.5439 | -36.6691
26.639 | -35.5443 | -36.6752
26.6451 | -35.5447 | -36.6809
26.6512 | -35.5452 | -36.6863
26.6573 | -35.5455 | -36.6916
26.6634 | -35.5459 | -36.6971
26.6695 | -35.5463 | -36.7026
26.6756 | -35.5467 | -36.7084
26.6817 | -35.5471 | -36.7137
26.6878 | -35.5474 | -36.7194
26.6939 | -35.5478 | -36.725
26.7 | -35.5482 | -36.7304
26.7061 | -35.5485 | -36.7359
26.7122 | -35.5489 | -36.7412
26.7183 | -35.5493 | -36.7465
26.7245 | -35.5497 | -36.7519
26.7306 | -35.55 | -36.7573
26.7367 | -35.5504 | -36.7631
26.7428 | -35.5508 | -36.7688
26.7489 | -35.5511 | -36.7744
26.755 | -35.5515 | -36.7797
26.7611 | -35.5519 | -36.7852
26.7672 | -35.5523 | -36.7908
26.7733 | -35.5527 | -36.7958
26.7794 | -35.553 | -36.8015
26.7855 | -35.5534 | -36.8069
26.7916 | -35.5538 | -36.8126
26.7977 | -35.5542 | -36.8183
26.8038 | -35.5546 | -36.8236
26.8099 | -35.555 | -36.8292
26.816 | -35.5554 | -36.835
26.8221 | -35.5557 | -36.8405
26.8282 | -35.5561 | -36.8456
26.8343 | -35.5564 | -36.8509
26.8404 | -35.5568 | -36.8563
26.8465 | -35.5572 | -36.8619
26.8526 | -35.5576 | -36.8673
26.8587 | -35.5579 | -36.8726
26.8649 | -35.5583 | -36.878
26.871 | -35.5587 | -36.8837
26.8771 | -35.559 | -36.8888
26.8832 | -35.5594 | -36.8945
26.8893 | -35.5598 | -36.8999
26.8954 | -35.5601 | -36.9054
26.9015 | -35.5604 | -36.911
26.9076 | -35.5608 | -36.9164
26.9137 | -35.5613 | -36.9217
26.9198 | -35.5616 | -36.9273
26.9259 | -35.5619 | -36.9326
26.932 | -35.5623 | -36.9385
26.9381 | -35.5627 | -36.9444
26.9442 | -35.5631 | -36.9499
26.9503 | -35.5635 | -36.9555
26.9564 | -35.5639 | -36.961
26.9625 | -35.5643 | -36.9664
26.9686 | -35.5646 | -36.9717
26.9747 | -35.565 | -36.9773
26.9808 | -35.5654 | -36.9828
26.9869 | -35.5657 | -36.9882
26.993 | -35.5661 | -36.9933
26.9991 | -35.5664 | -36.9986
27.0052 | -35.5668 | -37.004
27.0114 | -35.5672 | -37.0096
27.0175 | -35.5675 | -37.0148
27.0236 | -35.5679 | -37.0204
27.0297 | -35.5682 | -37.0256
27.0358 | -35.5686 | -37.0311
27.0419 | -35.569 | -37.0367
27.048 | -35.5693 | -37.0417
27.0541 | -35.5697 | -37.0467
27.0602 | -35.5701 | -37.052
27.0663 | -35.5705 | -37.0576
27.0724 | -35.5709 | -37.0628
27.0785 | -35.5713 | -37.0683
27.0846 | -35.5718 | -37.0733
27.0907 | -35.5721 | -37.0787
27.0968 | -35.5725 | -37.084
27.1029 | -35.5729 | -37.0894
27.109 | -35.5732 | -37.0947
27.1151 | -35.5736 | -37.1004
27.1212 | -35.574 | -37.1056
27.1273 | -35.5744 | -37.1112
27.1334 | -35.5748 | -37.117
27.1395 | -35.5752 | -37.1223
27.1456 | -35.5755 | -37.1278
27.1518 | -35.5759 | -37.133
27.1579 | -35.5763 | -37.1387
27.164 | -35.5766 | -37.1441
27.1701 | -35.577 | -37.1498
27.1762 | -35.5774 | -37.1552
27.1823 | -35.5777 | -37.1608
27.1884 | -35.5781 | -37.1663
27.1945 | -35.5784 | -37.1708
27.2006 | -35.5788 | -37.176
27.2067 | -35.5792 | -37.1812
27.2128 | -35.5795 | -37.1866
27.2189 | -35.5798 | -37.1918
27.225 | -35.5802 | -37.1971
27.2311 | -35.5807 | -37.2022
27.2372 | -35.581 | -37.2076
27.2433 | -35.5815 | -37.213
27.2494 | -35.5818 | -37.2185
27.2555 | -35.5822 | -37.224
27.2616 | -35.5826 | -37.2291
27.2677 | -35.583 | -37.2344
27.2738 | -35.5833 | -37.2398
27.2799 | -35.5838 | -37.2451
27.286 | -35.5841 | -37.2506
27.2921 | -35.5845 | -37.2563
27.2983 | -35.5849 | -37.2608
27.3044 | -35.5853 | -37.2662
27.3105 | -35.5857 | -37.2712
27.3166 | -35.586 | -37.2763
27.3227 | -35.5864 | -37.2818
27.3288 | -35.5868 | -37.2871
27.3349 | -35.5872 | -37.2919
27.341 | -35.5875 | -37.2969
27.3471 | -35.5878 | -37.3023
27.3532 | -35.5882 | -37.3077
27.3593 | -35.5886 | -37.3128
27.3654 | -35.5889 | -37.3182
27.3715 | -35.5893 | -37.3234
27.3776 | -35.5897 | -37.3287
27.3837 | -35.5901 | -37.3341
27.3898 | -35.5905 | -37.3394
27.3959 | -35.5909 | -37.3444
27.402 | -35.5912 | -37.3494
27.4081 | -35.5916 | -37.3547
27.4142 | -35.592 | -37.3602
27.4203 | -35.5923 | -37.3654
27.4264 | -35.5926 | -37.3707
27.4325 | -35.5929 | -37.3756
27.4387 | -35.5933 | -37.3807
27.4448 | -35.5936 | -37.386
27.4509 | -35.5939 | -37.3914
27.457 | -35.5943 | -37.3968
27.4631 | -35.5947 | -37.4022
27.4692 | -35.595 | -37.4079
27.4753 | -35.5954 | -37.4129
27.4814 | -35.5958 | -37.418
27.4875 | -35.5961 | -37.4231
27.4936 | -35.5965 | -37.4285
27.4997 | -35.5968 | -37.4334
27.5058 | -35.5972 | -37.4386
27.5119 | -35.5976 | -37.4435
27.518 | -35.5979 | -37.4489
27.5241 | -35.5983 | -37.4541
27.5302 | -35.5986 | -37.4593
27.5363 | -35.5991 | -37.4648
27.5424 | -35.5994 | -37.4693
27.5485 | -35.5998 | -37.4746
27.5546 | -35.6002 | -37.4797
27.5607 | -35.6006 | -37.4851
27.5668 | -35.601 | -37.4902
27.5729 | -35.6014 | -37.4953
27.5791 | -35.6018 | -37.5003
27.5852 | -35.6022 | -37.5057
27.5913 | -35.6025 | -37.511
27.5974 | -35.6029 | -37.5162
27.6035 | -35.6033 | -37.5214
27.6096 | -35.6037 | -37.5266
27.6157 | -35.604 | -37.532
27.6218 | -35.6044 | -37.5373
27.6279 | -35.6048 | -37.5424
27.634 | -35.6052 | -37.5477
27.6401 | -35.6055 | -37.5529
27.6462 | -35.6059 | -37.5581
27.6523 | -35.6063 | -37.5625
27.6584 | -35.6067 | -37.5676
27.6645 | -35.6071 | -37.5727
27.6706 | -35.6074 | -37.5781
27.6767 | -35.6078 | -37.5833
27.6828 | -35.6082 | -37.5883
27.6889 | -35.6086 | -37.5936
27.695 | -35.609 | -37.5987
27.7011 | -35.6094 | -37.6037
27.7072 | -35.6098 | -37.6089
27.7133 | -35.6102 | -37.6142
27.7194 | -35.6106 | -37.6191
27.7256 | -35.611 | -37.6243
27.7317 | -35.6113 | -37.6295
27.7378 | -35.6117 | -37.6347
27.7439 | -35.6121 | -37.6398
27.75 | -35.6125 | -37.645
27.7561 | -35.6128 | -37.65
27.7622 | -35.6132 | -37.6551
27.7683 | -35.6136 | -37.6604
27.7744 | -35.614 | -37.6658
27.7805 | -35.6144 | -37.6711
27.7866 | -35.6147 | -37.6762
27.7927 | -35.6151 | -37.6809
27.7988 | -35.6154 | -37.6858
27.8049 | -35.6158 | -37.6907
27.811 | -35.6161 | -37.6958
27.8171 | -35.6165 | -37.7007
27.8232 | -35.6169 | -37.7052
27.8293 | -35.6173 | -37.7103
27.8354 | -35.6177 | -37.7157
27.8415 | -35.6181 | -37.7207
27.8476 | -35.6184 | -37.7261
27.8537 | -35.6188 | -37.7314
27.8598 | -35.6191 | -37.736
27.866 | -35.6195 | -37.741
27.8721 | -35.6199 | -37.7461
27.8782 | -35.6203 | -37.7514
27.8843 | -35.6207 | -37.7562
27.8904 | -35.621 | -37.7613
27.8965 | -35.6214 | -37.766
27.9026 | -35.6218 | -37.7711
27.9087 | -35.6222 | -37.776
27.9148 | -35.6226 | -37.7812
27.9209 | -35.623 | -37.7862
27.927 | -35.6233 | -37.7908
27.9331 | -35.6237 | -37.7956
27.9392 | -35.6241 | -37.8009
27.9453 | -35.6245 | -37.8061
27.9514 | -35.6249 | -37.8112
27.9575 | -35.6252 | -37.8163
27.9636 | -35.6255 | -37.8214
27.9697 | -35.6259 | -37.8261
27.9758 | -35.6262 | -37.8317
27.9819 | -35.6266 | -37.8371
27.988 | -35.627 | -37.8422
27.9941 | -35.6273 | -37.8475
28.0002 | -35.6277 | -37.8524
28.0063 | -35.6281 | -37.8575
28.0125 | -35.6284 | -37.8627
28.0186 | -35.6288 | -37.8675
28.0247 | -35.6291 | -37.8727
28.0308 | -35.6295 | -37.8778
28.0369 | -35.6298 | -37.883
28.043 | -35.6301 | -37.8881
28.0491 | -35.6305 | -37.8931
28.0552 | -35.6308 | -37.8981
28.0613 | -35.6312 | -37.9028
28.0674 | -35.6316 | -37.9079
28.0735 | -35.632 | -37.9129

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4267 | -41.7846
21.4908 | -35.4824 | -34.952
16.4948 | -34.8591 | -27.6005
12.6602 | -34.4975 | -21.3556
9.7171 | -34.275 | -16.4739
7.4582 | -34.1286 | -12.5062
5.7244 | -34.0171 | -9.2941
4.3936 | -33.9576 | -6.6997
3.3722 | -33.9372 | -4.6595
2.5883 | -33.9254 | -3.1104
1.9866 | -33.9273 | -1.8748
1.5248 | -33.925 | -0.88008
1.1703 | -33.9324 | -0.23399
1 | -33.9416 | 0.085479
0.89822 | -33.9396 | 0.35838
0.68941 | -33.9185 | 0.93882
0.52919 | -33.9215 | 1.3872
0.40612 | -33.9155 | 1.9921
0.31171 | -33.9072 | 2.4448
0.23928 | -33.9131 | 3.1649
0.18358 | -33.918 | 3.5908
0.14091 | -33.9314 | 4.4705
0.10819 | -33.9536 | 5.6258
0.083017 | -33.9842 | 7.1018
0.063705 | -34.0196 | 8.9737
0.048923 | -34.0939 | 11.4671
0.037479 | -34.1977 | 14.6431
0.028801 | -34.3744 | 18.6623
0.022078 | -34.6834 | 24.0119
0.017023 | -35.1592 | 30.4149
0.013018 | -35.9652 | 38.0863
0.010014 | -37.0785 | 45.9772

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003727 | -8.6824
0.007454 | -26.8618
0.011182 | -34.3865
0.014909 | -39.0543
0.018636 | -41.4496
0.022363 | -43.668
0.02609 | -44.6666
0.029817 | -45.9172
0.033545 | -49.4886
0.037272 | -50.9354
0.040999 | -51.0336
0.044726 | -52.1766
0.048453 | -53.0074
0.05218 | -53.4959
0.055908 | -54.1913
0.059635 | -55.704
0.063362 | -56.3828
0.067089 | -57.449
0.070816 | -58.8638
0.074543 | -59.9508
0.078271 | -59.3562
0.081998 | -60.1082
0.085725 | -60.6249
0.089452 | -61.299
0.093179 | -61.9485
0.096906 | -62.4726
0.10063 | -62.833
0.10436 | -63.7871
0.10809 | -63.435
0.11181 | -63.3358
0.11554 | -64.0927
0.11927 | -65.3515
0.123 | -66.2826
0.12672 | -66.043
0.13045 | -66.2504
0.13418 | -66.1206
0.1379 | -65.7597
0.14163 | -65.698
0.14536 | -66.2698
0.14909 | -66.6857
0.15281 | -67.0706
0.15654 | -67.553
0.16027 | -68.3401
0.164 | -68.43
0.16772 | -69.8612
0.17145 | -70.2422
0.17518 | -69.8797
0.1789 | -68.9691
0.18263 | -69.4298
0.18636 | -69.7992
0.19009 | -71.0982
0.19381 | -71.3951
0.19754 | -71.91
0.20127 | -70.1428
0.20499 | -70.3517
0.20872 | -71.2267
0.21245 | -71.5866
0.21618 | -72.4822
0.2199 | -73.1125
0.22363 | -73.169
0.22736 | -73.1417
0.23109 | -72.6893
0.23481 | -72.8088
0.23854 | -72.5739
0.24227 | -72.4905
0.24599 | -72.7789
0.24972 | -72.41
0.25345 | -73.6187
0.25717 | -74.5514
0.2609 | -74.4131
0.26463 | -74.3513
0.26836 | -74.7418
0.27208 | -75.1964
0.27581 | -75.1825
0.27954 | -75.7255
0.28326 | -75.9129
0.28699 | -75.5855
0.29072 | -75.829
0.29445 | -75.3779
0.29817 | -75.4523
0.3019 | -75.5788
0.30563 | -75.6617
0.30935 | -75.9376
0.31308 | -75.5764
0.31681 | -75.8536
0.32054 | -75.7901
0.32426 | -76.6286
0.32799 | -76.2473
0.33172 | -76.8338
0.33544 | -77.1141
0.33917 | -76.3557
0.3429 | -77.5051
0.34663 | -77.6217
0.35035 | -77.1529
0.35408 | -77.0461
0.35781 | -77.0801
0.36154 | -77.5741
0.36526 | -77.3387
0.36899 | -77.8563
0.37272 | -78.1283
0.37644 | -77.0829
0.38017 | -76.6348
0.3839 | -77.2387
0.38763 | -77.958
0.39135 | -78.5885
0.39508 | -78.2054
0.39881 | -77.7153
0.40253 | -77.5709
0.40626 | -78.1993
0.40999 | -78.6364
0.41372 | -78.7208
0.41744 | -77.4034
0.42117 | -77.7006
0.4249 | -79.0474
0.42862 | -79.8963
0.43235 | -78.6951
0.43608 | -78.2354
0.43981 | -78.1306
0.44353 | -79.1019
0.44726 | -78.9924
0.45099 | -79.7299
0.45471 | -79.9942
0.45844 | -79.4363
0.46217 | -78.8363
0.4659 | -79.1952
0.46962 | -79.0922
0.47335 | -80.2728
0.47708 | -80.2549
0.4808 | -80.8197
0.48453 | -81.1039
0.48826 | -80.3784
0.49199 | -79.8804
0.49571 | -80.7882
0.49944 | -79.7026
0.50317 | -80.2972
0.50689 | -80.0193
0.51062 | -79.9623
0.51435 | -80.9626
0.51808 | -80.8348
0.5218 | -81.3599
0.52553 | -80.3953
0.52926 | -79.8866
0.53299 | -79.2082
0.53671 | -80.1914
0.54044 | -81.1078
0.54417 | -81.4578
0.54789 | -81.3514
0.55162 | -81.4991
0.55535 | -81.5625
0.55908 | -81.4777
0.5628 | -81.8064
0.56653 | -81.9902
0.57026 | -82.0409
0.57398 | -82.2374
0.57771 | -82.5255
0.58144 | -82.4342
0.58517 | -82.3387
0.58889 | -82.6829
0.59262 | -83.0601
0.59635 | -82.9185
0.60008 | -82.6061
0.6038 | -82.7895
0.60753 | -83.254
0.61126 | -83.2071
0.61498 | -82.7897
0.61871 | -82.6182
0.62244 | -82.6035
0.62616 | -82.2766
0.62989 | -82.3016
0.63362 | -83.1035
0.63735 | -83.9837
0.64107 | -84.0527
0.6448 | -83.6011
0.64853 | -83.5462
0.65226 | -83.274
0.65598 | -82.9915
0.65971 | -83.0816
0.66344 | -83.0947
0.66716 | -83.4322
0.67089 | -84.0593
0.67462 | -84.0988
0.67834 | -83.7795
0.68207 | -83.614
0.6858 | -83.8055
0.68953 | -83.7086
0.69325 | -83.3525
0.69698 | -83.4056
0.70071 | -83.7655
0.70444 | -84.0783
0.70816 | -84.3031
0.71189 | -84.4973
0.71562 | -84.5896
0.71934 | -84.5605
0.72307 | -84.5284
0.7268 | -84.6824
0.73053 | -84.4123
0.73425 | -84.093
0.73798 | -84.2414
0.74171 | -84.5873
0.74543 | -84.5642
0.74916 | -84.4809
0.75289 | -84.9385
0.75662 | -84.9909
0.76034 | -84.4036
0.76407 | -84.0702
0.7678 | -84.1734
0.77152 | -84.3605
0.77525 | -84.619
0.77898 | -85.426
0.78271 | -86.3574
0.78643 | -86.603
0.79016 | -86.4086
0.79389 | -86.3438
0.79761 | -86.447
0.80134 | -85.9288
0.80507 | -84.8946
0.8088 | -84.0647
0.81252 | -83.8857
0.81625 | -84.4349
0.81998 | -85.1761
0.82371 | -85.4464
0.82743 | -85.6673
0.83116 | -86.0791
0.83489 | -86.2981
0.83861 | -85.8374
0.84234 | -85.4211
0.84607 | -85.7084
0.84979 | -85.8818
0.85352 | -85.8069
0.85725 | -85.8656
0.86098 | -85.84
0.8647 | -86.0544
0.86843 | -86.3533
0.87216 | -86.685
0.87589 | -86.7704
0.87961 | -86.4922
0.88334 | -86.213
0.88707 | -86.3327
0.89079 | -86.5135
0.89452 | -86.1013
0.89825 | -85.9571
0.90197 | -86.3647
0.9057 | -86.6017
0.90943 | -86.5392
0.91316 | -86.5259
0.91688 | -86.3496
0.92061 | -86.2673
0.92434 | -86.4359
0.92807 | -86.482
0.93179 | -86.4912
0.93552 | -86.8131
0.93925 | -87.0038
0.94297 | -87.1197
0.9467 | -87.3419
0.95043 | -87.599
0.95416 | -87.3428
0.95788 | -86.8658
0.96161 | -86.4536
0.96534 | -86.0658
0.96906 | -86.0179
0.97279 | -86.581
0.97652 | -87.1371
0.98025 | -87.1793
0.98397 | -86.8822
0.9877 | -86.7591
0.99143 | -86.8976
0.99516 | -87.1261
0.99888 | -86.8577
1.0026 | -86.6385
1.0063 | -87.0685
1.0101 | -87.2728
1.0138 | -87.3765
1.0175 | -87.5093
1.0212 | -87.4419
1.025 | -87.6284
1.0287 | -87.1628
1.0324 | -87.4422
1.0362 | -87.9942
1.0399 | -88.01
1.0436 | -87.5803
1.0473 | -87.3511
1.0511 | -87.3648
1.0548 | -87.4655
1.0585 | -87.5123
1.0622 | -87.3424
1.066 | -87.3875
1.0697 | -87.56
1.0734 | -87.7341
1.0772 | -87.9385
1.0809 | -88.0861
1.0846 | -88.1271
1.0883 | -88.0751
1.0921 | -87.8255
1.0958 | -87.4591
1.0995 | -87.2612
1.1032 | -87.3429
1.107 | -87.6456
1.1107 | -87.9719
1.1144 | -88.1419
1.1182 | -88.198
1.1219 | -88.2317
1.1256 | -88.263
1.1293 | -88.3026
1.1331 | -88.2661
1.1368 | -88.069
1.1405 | -87.8819
1.1442 | -87.9145
1.148 | -88.1386
1.1517 | -88.4394
1.1554 | -88.7801
1.1591 | -89.0798
1.1629 | -89.1696
1.1666 | -89.0441
1.1703 | -88.8118
1.1741 | -88.6411
1.1778 | -88.6578
1.1815 | -88.8531
1.1852 | -89.2178
1.189 | -89.6182
1.1927 | -89.7271
1.1964 | -89.471
1.2001 | -89.1371
1.2039 | -88.9742
1.2076 | -88.9983
1.2113 | -89.1003
1.2151 | -89.1245
1.2188 | -89.0148
1.2225 | -88.9449
1.2262 | -88.9329
1.23 | -88.8452
1.2337 | -88.7345
1.2374 | -88.7309
1.2411 | -88.9655
1.2449 | -89.4639
1.2486 | -89.8616
1.2523 | -89.8917
1.2561 | -89.8062
1.2598 | -89.7709
1.2635 | -89.7249
1.2672 | -89.7335
1.271 | -89.878
1.2747 | -90.0623
1.2784 | -90.1171
1.2821 | -90.0412
1.2859 | -89.9324
1.2896 | -89.8374
1.2933 | -89.6843
1.2971 | -89.5118
1.3008 | -89.5604
1.3045 | -89.8656
1.3082 | -90.0771
1.312 | -89.9882
1.3157 | -89.826
1.3194 | -89.7775
1.3231 | -89.8116
1.3269 | -89.9328
1.3306 | -90.0776
1.3343 | -89.9599
1.3381 | -89.589
1.3418 | -89.3869
1.3455 | -89.4992
1.3492 | -89.6988
1.353 | -89.7925
1.3567 | -89.7872
1.3604 | -89.8771
1.3641 | -90.2764
1.3679 | -90.8221
1.3716 | -91.0817
1.3753 | -90.9881
1.3791 | -90.8038
1.3828 | -90.5938
1.3865 | -90.334
1.3902 | -90.1394
1.394 | -90.1416
1.3977 | -90.2478
1.4014 | -90.2559
1.4051 | -90.0881
1.4089 | -89.877
1.4126 | -89.7926
1.4163 | -89.8173
1.4201 | -89.8364
1.4238 | -89.9444
1.4275 | -90.2139
1.4312 | -90.4241
1.435 | -90.518
1.4387 | -90.7359
1.4424 | -90.9722
1.4461 | -90.922
1.4499 | -90.6681
1.4536 | -90.4576
1.4573 | -90.3728
1.4611 | -90.3723
1.4648 | -90.4417
1.4685 | -90.542
1.4722 | -90.5908
1.476 | -90.6303
1.4797 | -90.6819
1.4834 | -90.7165
1.4871 | -90.7772
1.4909 | -90.8184
1.4946 | -90.7804
1.4983 | -90.7679
1.5021 | -90.8577
1.5058 | -90.9538
1.5095 | -90.9864
1.5132 | -90.9917
1.517 | -91.1195
1.5207 | -91.3972
1.5244 | -91.4875
1.5281 | -91.243
1.5319 | -91.028
1.5356 | -91.0182
1.5393 | -91.0612
1.543 | -91.0084
1.5468 | -90.9298
1.5505 | -90.9321
1.5542 | -90.9686
1.558 | -90.9397
1.5617 | -90.796
1.5654 | -90.671
1.5691 | -90.7648
1.5729 | -91.0314
1.5766 | -91.212
1.5803 | -91.1882
1.584 | -91.1448
1.5878 | -91.237
1.5915 | -91.37
1.5952 | -91.5006
1.599 | -91.6701
1.6027 | -91.8275
1.6064 | -91.8524
1.6101 | -91.6272
1.6139 | -91.4697
1.6176 | -91.3752
1.6213 | -91.3701
1.625 | -91.3963
1.6288 | -91.352
1.6325 | -91.2045
1.6362 | -91.0406
1.64 | -90.9856
1.6437 | -91.0916
1.6474 | -91.3124
1.6511 | -91.5275
1.6549 | -91.7109
1.6586 | -91.9374
1.6623 | -92.2505
1.666 | -92.5528
1.6698 | -92.7434
1.6735 | -92.8455
1.6772 | -92.9295
1.681 | -92.9612
1.6847 | -92.8456
1.6884 | -92.5907
1.6921 | -92.348
1.6959 | -92.2357
1.6996 | -92.2125
1.7033 | -92.1722
1.707 | -92.106
1.7108 | -92.0883
1.7145 | -92.1779
1.7182 | -92.3296
1.722 | -92.4735
1.7257 | -92.6199
1.7294 | -92.7636
1.7331 | -92.8546
1.7369 | -92.8397
1.7406 | -92.7425
1.7443 | -92.6312
1.748 | -92.5399
1.7518 | -92.4616
1.7555 | -92.362
1.7592 | -92.2463
1.763 | -92.1617
1.7667 | -92.1322
1.7704 | -92.1478
1.7741 | -92.1733
1.7779 | -92.1478
1.7816 | -92.0604
1.7853 | -91.9939
1.789 | -92.0513
1.7928 | -92.1941
1.7965 | -92.2859
1.8002 | -92.2348
1.804 | -92.1356
1.8077 | -92.1121
1.8114 | -92.1856
1.8151 | -92.3119
1.8189 | -92.4889
1.8226 | -92.7064
1.8263 | -92.9135
1.83 | -93.0411
1.8338 | -93.1114
1.8375 | -93.1867
1.8412 | -93.2652
1.845 | -93.3058
1.8487 | -93.2979
1.8524 | -93.2968
1.8561 | -93.3505
1.8599 | -93.4407
1.8636 | -93.5193
1.8673 | -93.5271
1.871 | -93.445
1.8748 | -93.3369
1.8785 | -93.2533
1.8822 | -93.1701
1.8859 | -93.0983
1.8897 | -93.0373
1.8934 | -92.9582
1.8971 | -92.8805
1.9009 | -92.8938
1.9046 | -92.969
1.9083 | -92.985
1.912 | -92.9056
1.9158 | -92.8406
1.9195 | -92.8574
1.9232 | -92.9711
1.9269 | -93.1253
1.9307 | -93.2109
1.9344 | -93.17
1.9381 | -93.0191
1.9419 | -92.7886
1.9456 | -92.5738
1.9493 | -92.5026
1.953 | -92.6116
1.9568 | -92.8199
1.9605 | -93.0162
1.9642 | -93.1393
1.9679 | -93.1865
1.9717 | -93.2097
1.9754 | -93.2233
1.9791 | -93.2136
1.9829 | -93.2096
1.9866 | -93.2746
1.9903 | -93.3919
1.994 | -93.4741
1.9978 | -93.4632
2.0015 | -93.3937
2.0052 | -93.3491
2.0089 | -93.375
2.0127 | -93.432
2.0164 | -93.4442
2.0201 | -93.3926
2.0239 | -93.3277
2.0276 | -93.3104
2.0313 | -93.3769
2.035 | -93.5116
2.0388 | -93.6865
2.0425 | -93.8866
2.0462 | -94.1026
2.0499 | -94.2632
2.0537 | -94.2959
2.0574 | -94.2381
2.0611 | -94.2032
2.0649 | -94.2145
2.0686 | -94.2025
2.0723 | -94.1174
2.076 | -93.9893
2.0798 | -93.8644
2.0835 | -93.7569
2.0872 | -93.6454
2.0909 | -93.5216
2.0947 | -93.3857
2.0984 | -93.3008
2.1021 | -93.3616
2.1059 | -93.6013
2.1096 | -94.0007
2.1133 | -94.471
2.117 | -94.8896
2.1208 | -95.1265
2.1245 | -95.1237
2.1282 | -94.933
2.1319 | -94.6887
2.1357 | -94.5432
2.1394 | -94.5753
2.1431 | -94.7313
2.1469 | -94.8887
2.1506 | -94.8023
2.1543 | -94.7034
2.158 | -94.5382
2.1618 | -94.3546
2.1655 | -94.207
2.1692 | -94.1389
2.1729 | -94.1543
2.1767 | -94.2273
2.1804 | -94.3143
2.1841 | -94.3979
2.1878 | -94.4636
2.1916 | -94.4912
2.1953 | -94.456
2.199 | -94.3884
2.2028 | -94.3594
2.2065 | -94.4377
2.2102 | -94.6018
2.2139 | -94.7756
2.2177 | -94.8743
2.2214 | -94.853
2.2251 | -94.715
2.2288 | -94.515
2.2326 | -94.2985
2.2363 | -94.1316
2.24 | -94.0251
2.2438 | -93.9755
2.2475 | -93.9673
2.2512 | -93.9941
2.2549 | -94.0487
2.2587 | -94.1162
2.2624 | -94.1982
2.2661 | -94.3075
2.2698 | -94.4379
2.2736 | -94.5671
2.2773 | -94.6395
2.281 | -94.6266
2.2848 | -94.5224
2.2885 | -94.3798
2.2922 | -94.2573
2.2959 | -94.1861
2.2997 | -94.1609
2.3034 | -94.1592
2.3071 | -94.1616
2.3108 | -94.1823
2.3146 | -94.2535
2.3183 | -94.4065
2.322 | -94.6604
2.3258 | -95.0003
2.3295 | -95.3662
2.3332 | -95.6926
2.3369 | -95.8643
2.3407 | -95.7936
2.3444 | -95.4864
2.3481 | -95.0683
2.3518 | -94.675
2.3556 | -94.4087
2.3593 | -94.3117
2.363 | -94.3697
2.3668 | -94.538
2.3705 | -94.7224
2.3742 | -94.8312
2.3779 | -94.8399
2.3817 | -94.7717
2.3854 | -94.682
2.3891 | -94.6286
2.3928 | -94.6257
2.3966 | -94.6667
2.4003 | -94.7274
2.404 | -94.765
2.4078 | -94.7574
2.4115 | -94.7101
2.4152 | -94.6588
2.4189 | -94.6626
2.4227 | -94.7553
2.4264 | -94.9147
2.4301 | -95.0776
2.4338 | -95.2233
2.4376 | -95.3397
2.4413 | -95.4172
2.445 | -95.4404
2.4488 | -95.3992
2.4525 | -95.3032
2.4562 | -95.1791
2.4599 | -95.0468
2.4637 | -94.9035
2.4674 | -94.7526
2.4711 | -94.5819
2.4748 | -94.3848
2.4786 | -94.2198
2.4823 | -94.1586
2.486 | -94.2281
2.4897 | -94.3758
2.4935 | -94.5416
2.4972 | -94.6558
2.5009 | -94.7023
2.5047 | -94.7118
2.5084 | -94.7196
2.5121 | -94.7607
2.5158 | -94.8237
2.5196 | -94.8595
2.5233 | -94.8499
2.527 | -94.805
2.5307 | -94.7873
2.5345 | -94.8299
2.5382 | -94.97
2.5419 | -95.1894
2.5457 | -95.432
2.5494 | -95.6299
2.5531 | -95.749
2.5568 | -95.792
2.5606 | -95.7764
2.5643 | -95.7198
2.568 | -95.6309
2.5717 | -95.4883
2.5755 | -95.3064
2.5792 | -95.0937
2.5829 | -94.89
2.5867 | -94.6996
2.5904 | -94.5239
2.5941 | -94.3747
2.5978 | -94.2627
2.6016 | -94.2035
2.6053 | -94.2132
2.609 | -94.3167
2.6127 | -94.5055
2.6165 | -94.7283
2.6202 | -94.9163
2.6239 | -95.0153
2.6277 | -95.0265
2.6314 | -94.976
2.6351 | -94.9179
2.6388 | -94.9085
2.6426 | -94.9646
2.6463 | -95.0659
2.65 | -95.1789
2.6537 | -95.2952
2.6575 | -95.396
2.6612 | -95.4712
2.6649 | -95.5181
2.6687 | -95.5415
2.6724 | -95.5441
2.6761 | -95.5032
2.6798 | -95.3928
2.6836 | -95.2108
2.6873 | -95.0299
2.691 | -94.9225
2.6947 | -94.8967
2.6985 | -94.9544
2.7022 | -95.0796
2.7059 | -95.2198
2.7097 | -95.3358
2.7134 | -95.4193
2.7171 | -95.4754
2.7208 | -95.5349
2.7246 | -95.6124
2.7283 | -95.7184
2.732 | -95.8545
2.7357 | -95.975
2.7395 | -96.0412
2.7432 | -96.0521
2.7469 | -96.0263
2.7507 | -95.9793
2.7544 | -95.9372
2.7581 | -95.8935
2.7618 | -95.8559
2.7656 | -95.8296
2.7693 | -95.819
2.773 | -95.852
2.7767 | -95.9137
2.7805 | -95.9914
2.7842 | -96.0682
2.7879 | -96.1295
2.7917 | -96.1612
2.7954 | -96.1563
2.7991 | -96.1322
2.8028 | -96.0866
2.8066 | -96.0449
2.8103 | -96.021
2.814 | -96.0168
2.8177 | -96.0238
2.8215 | -96.0199
2.8252 | -95.9984
2.8289 | -95.9506
2.8327 | -95.9031
2.8364 | -95.8721
2.8401 | -95.8684
2.8438 | -95.9003
2.8476 | -95.9311
2.8513 | -95.974
2.855 | -96.0093
2.8587 | -96.0345
2.8625 | -96.0753
2.8662 | -96.1253
2.8699 | -96.1777
2.8736 | -96.233
2.8774 | -96.2899
2.8811 | -96.3413
2.8848 | -96.3959
2.8886 | -96.446
2.8923 | -96.4797
2.896 | -96.4917
2.8997 | -96.4554
2.9035 | -96.4013
2.9072 | -96.3504
2.9109 | -96.2965
2.9146 | -96.2495
2.9184 | -96.1887
2.9221 | -96.0984
2.9258 | -95.9737
2.9296 | -95.8476
2.9333 | -95.76
2.937 | -95.7408
2.9407 | -95.7987
2.9445 | -95.932
2.9482 | -96.1147
2.9519 | -96.2965
2.9556 | -96.4439
2.9594 | -96.5356
2.9631 | -96.5723
2.9668 | -96.5688
2.9706 | -96.5505
2.9743 | -96.5153
2.978 | -96.3896
2.9817 | -96.5007
2.9855 | -96.444
2.9892 | -96.3616
2.9929 | -96.2576
2.9966 | -96.1522
3.0004 | -96.0686
3.0041 | -96.043
3.0078 | -96.0909
3.0116 | -96.2004
3.0153 | -96.3454
3.019 | -96.4611
3.0227 | -96.4833
3.0265 | -96.6433
3.0302 | -96.6706
3.0339 | -96.6471
3.0376 | -96.5893
3.0414 | -96.5293
3.0451 | -96.4954
3.0488 | -96.5032
3.0526 | -96.5553
3.0563 | -96.637
3.06 | -96.7087
3.0637 | -96.7368
3.0675 | -96.7018
3.0712 | -96.6208
3.0749 | -96.5356
3.0786 | -96.49
3.0824 | -96.4923
3.0861 | -96.5402
3.0898 | -96.6091
3.0936 | -96.6671
3.0973 | -96.7091
3.101 | -96.7484
3.1047 | -96.7806
3.1085 | -96.8112
3.1122 | -96.8301
3.1159 | -96.8454
3.1196 | -96.8454
3.1234 | -96.8285
3.1271 | -96.7919
3.1308 | -96.7447
3.1346 | -96.693
3.1383 | -96.649
3.142 | -96.6386
3.1457 | -96.6576
3.1495 | -96.6836
3.1532 | -96.7133
3.1569 | -96.7516
3.1606 | -96.8097
3.1644 | -96.8767
3.1681 | -96.9331
3.1718 | -96.9341
3.1755 | -96.8679
3.1793 | -96.7692
3.183 | -96.6841
3.1867 | -96.6518
3.1905 | -96.6723
3.1942 | -96.7316
3.1979 | -96.8153
3.2016 | -96.8906
3.2054 | -96.9409
3.2091 | -96.9801
3.2128 | -97.0216
3.2165 | -97.0924
3.2203 | -97.1983
3.224 | -97.3125
3.2277 | -97.3998
3.2315 | -97.4668
3.2352 | -97.4912
3.2389 | -97.4871
3.2426 | -97.4717
3.2464 | -97.4507
3.2501 | -97.4399
3.2538 | -97.4504
3.2575 | -97.4872
3.2613 | -97.5365
3.265 | -97.5901
3.2687 | -97.6181
3.2725 | -97.6269
3.2762 | -97.604
3.2799 | -97.5582
3.2836 | -97.4967
3.2874 | -97.4287
3.2911 | -97.385
3.2948 | -97.3662
3.2985 | -97.3731
3.3023 | -97.3995
3.306 | -97.4302
3.3097 | -97.4621
3.3135 | -97.4844
3.3172 | -97.5112
3.3209 | -97.5242
3.3246 | -97.5327
3.3284 | -97.5227
3.3321 | -97.4929
3.3358 | -97.4432
3.3395 | -97.3806
3.3433 | -97.326
3.347 | -97.2918
3.3507 | -97.2748
3.3545 | -97.2839
3.3582 | -97.3087
3.3619 | -97.3404
3.3656 | -97.3712
3.3694 | -97.3981
3.3731 | -97.4163
3.3768 | -97.4255
3.3805 | -97.418
3.3843 | -97.3961
3.388 | -97.3797
3.3917 | -97.371
3.3955 | -97.3771
3.3992 | -97.3993
3.4029 | -97.4361
3.4066 | -97.4963
3.4104 | -97.5366
3.4141 | -97.5526
3.4178 | -97.5396
3.4215 | -97.5058
3.4253 | -97.4696
3.429 | -97.4351
3.4327 | -97.4114
3.4365 | -97.382
3.4402 | -97.339
3.4439 | -97.2613
3.4476 | -97.1618
3.4514 | -97.0586
3.4551 | -96.9649
3.4588 | -96.9036
3.4625 | -96.8924
3.4663 | -96.9565
3.47 | -97.0716
3.4737 | -97.2122
3.4775 | -97.355
3.4812 | -97.4734
3.4849 | -97.5466
3.4886 | -97.5685
3.4924 | -97.5574
3.4961 | -97.527
3.4998 | -97.4875
3.5035 | -97.4375
3.5073 | -97.3983
3.511 | -97.394
3.5147 | -97.4235
3.5184 | -97.4852
3.5222 | -97.5539
3.5259 | -97.6082
3.5296 | -97.6387
3.5334 | -97.6467
3.5371 | -97.6466
3.5408 | -97.6822
3.5445 | -97.7451
3.5483 | -97.8178
3.552 | -97.8919
3.5557 | -97.9459
3.5594 | -97.97
3.5632 | -97.9617
3.5669 | -97.9355
3.5706 | -97.9191
3.5744 | -97.9134
3.5781 | -97.9128
3.5818 | -97.9156
3.5855 | -97.893
3.5893 | -97.8454
3.593 | -97.7826
3.5967 | -97.7362
3.6004 | -97.7379
3.6042 | -97.7748
3.6079 | -97.8262
3.6116 | -97.89
3.6154 | -97.9556
3.6191 | -98.0339
3.6228 | -98.1217
3.6265 | -98.2078
3.6303 | -98.2771
3.634 | -98.3093
3.6377 | -98.2779
3.6414 | -98.2127
3.6452 | -98.1307
3.6489 | -98.048
3.6526 | -97.9718
3.6564 | -97.9255
3.6601 | -97.9299
3.6638 | -97.9736
3.6675 | -98.0394
3.6713 | -98.1169
3.675 | -98.1815
3.6787 | -98.2129
3.6824 | -98.1958
3.6862 | -98.1343
3.6899 | -98.0466
3.6936 | -97.9581
3.6974 | -97.8809
3.7011 | -97.8294
3.7048 | -97.814
3.7085 | -97.8155
3.7123 | -97.8145
3.716 | -97.784
3.7197 | -97.7257
3.7234 | -97.6444
3.7272 | -97.5445
3.7309 | -97.4565
3.7346 | -97.4079
3.7384 | -97.4092
3.7421 | -97.4685
3.7458 | -97.5845
3.7495 | -97.7384
3.7533 | -97.8983
3.757 | -98.0341
3.7607 | -98.1016
3.7644 | -98.1877
3.7682 | -98.2442
3.7719 | -98.2751
3.7756 | -98.2914
3.7794 | -98.2816
3.7831 | -98.2427
3.7868 | -98.1916
3.7905 | -98.1317
3.7943 | -98.0849
3.798 | -98.0585
3.8017 | -98.0524
3.8054 | -98.0608
3.8092 | -98.0835
3.8129 | -98.1158
3.8166 | -98.1513
3.8203 | -98.1858
3.8241 | -98.2213
3.8278 | -98.2604
3.8315 | -98.2987
3.8353 | -98.3193
3.839 | -98.3192
3.8427 | -98.3134
3.8464 | -98.3158
3.8502 | -98.333
3.8539 | -98.3762
3.8576 | -98.4308
3.8613 | -98.4943
3.8651 | -98.5556
3.8688 | -98.6005
3.8725 | -98.6334
3.8763 | -98.6631
3.88 | -98.691
3.8837 | -98.7164
3.8874 | -98.7545
3.8912 | -98.8
3.8949 | -98.8521
3.8986 | -98.9067
3.9023 | -98.9579
3.9061 | -99.0011
3.9098 | -99.0405
3.9135 | -99.0812
3.9173 | -99.1229
3.921 | -99.1758
3.9247 | -99.2295
3.9284 | -99.2834
3.9322 | -99.3434
3.9359 | -99.3893
3.9396 | -99.4219
3.9433 | -99.4365
3.9471 | -99.4388
3.9508 | -99.4295
3.9545 | -99.4105
3.9583 | -99.4028
3.962 | -99.3881
3.9657 | -99.3638
3.9694 | -99.3257
3.9732 | -99.2643
3.9769 | -99.1937
3.9806 | -99.1127
3.9843 | -99.0312
3.9881 | -98.9695
3.9918 | -98.9209
3.9955 | -98.8873
3.9993 | -98.8865
4.003 | -98.9167
4.0067 | -98.9661
4.0104 | -99.0086
4.0142 | -99.0443
4.0179 | -99.0767
4.0216 | -99.0931
4.0253 | -99.1047
4.0291 | -99.1063
4.0328 | -99.1132
4.0365 | -99.1116
4.0403 | -99.1036
4.044 | -99.1063
4.0477 | -99.1061
4.0514 | -99.0987
4.0552 | -99.1047
4.0589 | -99.1143
4.0626 | -99.1186
4.0663 | -99.1248
4.0701 | -99.1331
4.0738 | -99.1546
4.0775 | -99.1855
4.0813 | -99.2327
4.085 | -99.2951
4.0887 | -99.3723
4.0924 | -99.4568
4.0962 | -99.5302
4.0999 | -99.5831
4.1036 | -99.6006
4.1073 | -99.5805
4.1111 | -99.5403
4.1148 | -99.4827
4.1185 | -99.4052
4.1223 | -99.3268
4.126 | -99.2535
4.1297 | -99.1762
4.1334 | -99.1108
4.1372 | -99.0676
4.1409 | -99.0532
4.1446 | -99.072
4.1483 | -99.1293
4.1521 | -99.2141
4.1558 | -99.3124
4.1595 | -99.4022
4.1632 | -99.4755
4.167 | -99.5104
4.1707 | -99.5057
4.1744 | -99.4674
4.1782 | -99.4075
4.1819 | -99.3352
4.1856 | -99.2603
4.1893 | -99.1932
4.1931 | -99.1488
4.1968 | -99.1254
4.2005 | -99.1177
4.2042 | -99.1187
4.208 | -99.1384
4.2117 | -99.1606
4.2154 | -99.1746
4.2192 | -99.1731
4.2229 | -99.1594
4.2266 | -99.1388
4.2303 | -99.1196
4.2341 | -99.1127
4.2378 | -99.1176
4.2415 | -99.1329
4.2452 | -99.1677
4.249 | -99.2246
4.2527 | -99.2991
4.2564 | -99.3771
4.2602 | -99.4657
4.2639 | -99.5559
4.2676 | -99.6509
4.2713 | -99.7221
4.2751 | -99.7765
4.2788 | -99.8138
4.2825 | -99.8393
4.2862 | -99.8538
4.29 | -99.8555
4.2937 | -99.8594
4.2974 | -99.8683
4.3012 | -99.9114
4.3049 | -99.9473
4.3086 | -99.9888
4.3123 | -100.0326
4.3161 | -100.0766
4.3198 | -100.116
4.3235 | -100.1441
4.3272 | -100.1382
4.331 | -100.1118
4.3347 | -100.0605
4.3384 | -99.9791
4.3422 | -99.8754
4.3459 | -99.7551
4.3496 | -99.6294
4.3533 | -99.5126
4.3571 | -99.4183
4.3608 | -99.3408
4.3645 | -99.2851
4.3682 | -99.2532
4.372 | -99.2458
4.3757 | -99.2641
4.3794 | -99.3092
4.3832 | -99.3576
4.3869 | -99.4054
4.3906 | -99.4608
4.3943 | -99.5079
4.3981 | -99.5292
4.4018 | -99.5209
4.4055 | -99.4861
4.4092 | -99.4462
4.413 | -99.4081
4.4167 | -99.3722
4.4204 | -99.3386
4.4242 | -99.3059
4.4279 | -99.271
4.4316 | -99.2446
4.4353 | -99.2227
4.4391 | -99.2148
4.4428 | -99.2236
4.4465 | -99.2544
4.4502 | -99.3043
4.454 | -99.3834
4.4577 | -99.4807
4.4614 | -99.5735
4.4652 | -99.6606
4.4689 | -99.7324
4.4726 | -99.7696
4.4763 | -99.7801
4.4801 | -99.7584
4.4838 | -99.7067
4.4875 | -99.6422
4.4912 | -99.5786
4.495 | -99.5318
4.4987 | -99.5094
4.5024 | -99.5153
4.5061 | -99.5577
4.5099 | -99.63
4.5136 | -99.7303
4.5173 | -99.8474
4.5211 | -99.9691
4.5248 | -100.0768
4.5285 | -100.1708
4.5322 | -100.2509
4.536 | -100.3149
4.5397 | -100.3628
4.5434 | -100.3946
4.5471 | -100.3981
4.5509 | -100.3716
4.5546 | -100.3231
4.5583 | -100.2578
4.5621 | -100.165
4.5658 | -100.0532
4.5695 | -99.9331
4.5732 | -99.8182
4.577 | -99.7137
4.5807 | -99.6248
4.5844 | -99.5482
4.5881 | -99.4936
4.5919 | -99.4682
4.5956 | -99.4653
4.5993 | -99.4743
4.6031 | -99.4955
4.6068 | -99.5217
4.6105 | -99.5464
4.6142 | -99.5673
4.618 | -99.5863
4.6217 | -99.6066
4.6254 | -99.6224
4.6291 | -99.6467
4.6329 | -99.6895
4.6366 | -99.7439
4.6403 | -99.8012
4.6441 | -99.8492
4.6478 | -99.8933
4.6515 | -99.9267
4.6552 | -99.9404
4.659 | -99.939
4.6627 | -99.9217
4.6664 | -99.8926
4.6701 | -99.8653
4.6739 | -99.8344
4.6776 | -99.8145
4.6813 | -99.806
4.6851 | -99.813
4.6888 | -99.8349
4.6925 | -99.8747
4.6962 | -99.9283
4.7 | -99.9849
4.7037 | -100.0436
4.7074 | -100.1042
4.7111 | -100.172
4.7149 | -100.2429
4.7186 | -100.3071
4.7223 | -100.3736
4.7261 | -100.4471
4.7298 | -100.5149
4.7335 | -100.576
4.7372 | -100.6334
4.741 | -100.6747
4.7447 | -100.7068
4.7484 | -100.7167
4.7521 | -100.7115
4.7559 | -100.7
4.7596 | -100.6868
4.7633 | -100.6833
4.7671 | -100.6872
4.7708 | -100.6939
4.7745 | -100.7081
4.7782 | -100.7121
4.782 | -100.6949
4.7857 | -100.6568
4.7894 | -100.6013
4.7931 | -100.5338
4.7969 | -100.4692
4.8006 | -100.4081
4.8043 | -100.3666
4.8081 | -100.3461
4.8118 | -100.3498
4.8155 | -100.3762
4.8192 | -100.4215
4.823 | -100.4761
4.8267 | -100.5322
4.8304 | -100.5805
4.8341 | -100.6054
4.8379 | -100.5991
4.8416 | -100.6078
4.8453 | -100.6062
4.8491 | -100.5915
4.8528 | -100.5782
4.8565 | -100.5653
4.8602 | -100.5539
4.864 | -100.5421
4.8677 | -100.5333
4.8714 | -100.5225
4.8751 | -100.5222
4.8789 | -100.5254
4.8826 | -100.5362
4.8863 | -100.5468
4.89 | -100.5533
4.8938 | -100.5549
4.8975 | -100.5507
4.9012 | -100.5437
4.905 | -100.5444
4.9087 | -100.5671
4.9124 | -100.603
4.9161 | -100.6546
4.9199 | -100.7237
4.9236 | -100.7991
4.9273 | -100.8682
4.931 | -100.9189
4.9348 | -100.9486
4.9385 | -100.9581
4.9422 | -100.9573
4.946 | -100.9398
4.9497 | -100.9161
4.9534 | -100.8862
4.9571 | -100.8591
4.9609 | -100.7886
4.9646 | -100.6947
4.9683 | -100.6016
4.972 | -100.4815
4.9758 | -100.3427
4.9795 | -100.1831
4.9832 | -100.0123
4.987 | -99.8469
4.9907 | -99.7086
4.9944 | -99.6105
4.9981 | -99.5666
5.0019 | -99.5807
5.0056 | -99.6512
5.0093 | -99.7701
5.013 | -99.9267
5.0168 | -100.1015
5.0205 | -100.2781
5.0242 | -100.44
5.028 | -100.5734
5.0317 | -100.678
5.0354 | -100.7468
5.0391 | -100.807
5.0429 | -100.8379
5.0466 | -100.8293
5.0503 | -100.8164
5.054 | -100.7996
5.0578 | -100.7889
5.0615 | -100.7811
5.0652 | -100.7786
5.069 | -100.7804
5.0727 | -100.7759
5.0764 | -100.7568
5.0801 | -100.7365
5.0839 | -100.7104
5.0876 | -100.6897
5.0913 | -100.6764
5.095 | -100.6785
5.0988 | -100.6958
5.1025 | -100.735
5.1062 | -100.7819
5.11 | -100.8283
5.1137 | -100.8822
5.1174 | -100.9284
5.1211 | -100.9633
5.1249 | -100.9878
5.1286 | -100.9933
5.1323 | -100.9808
5.136 | -100.9506
5.1398 | -100.908
5.1435 | -100.8576
5.1472 | -100.802
5.1509 | -100.7593
5.1547 | -100.7355
5.1584 | -100.726
5.1621 | -100.7245
5.1659 | -100.7308
5.1696 | -100.7562
5.1733 | -100.7961
5.177 | -100.8402
5.1808 | -100.8846
5.1845 | -100.9313
5.1882 | -100.9581
5.1919 | -100.9677
5.1957 | -100.9773
5.1994 | -100.9728
5.2031 | -100.9593
5.2069 | -100.9442
5.2106 | -100.9277
5.2143 | -100.9128
5.218 | -100.9049
5.2218 | -100.9008
5.2255 | -100.9097
5.2292 | -100.9188
5.2329 | -100.9344
5.2367 | -100.9522
5.2404 | -100.9729
5.2441 | -100.9886
5.2479 | -100.975
5.2516 | -100.9528
5.2553 | -100.9365
5.259 | -100.9192
5.2628 | -100.8935
5.2665 | -100.8853
5.2702 | -100.8861
5.2739 | -100.8889
5.2777 | -100.8983
5.2814 | -100.9054
5.2851 | -100.9143
5.2889 | -100.9165
5.2926 | -100.919
5.2963 | -100.9227
5.3 | -100.9199
5.3038 | -100.9182
5.3075 | -100.9296
5.3112 | -100.9448
5.3149 | -100.9725
5.3187 | -101.0136
5.3224 | -101.0604
5.3261 | -101.1297
5.3299 | -101.1976
5.3336 | -101.2538
5.3373 | -101.3026
5.341 | -101.3456
5.3448 | -101.3703
5.3485 | -101.3771
5.3522 | -101.3717
5.3559 | -101.3538
5.3597 | -101.3357
5.3634 | -101.3053
5.3671 | -101.2683
5.3709 | -101.2329
5.3746 | -101.2051
5.3783 | -101.1673
5.382 | -101.1386
5.3858 | -101.1023
5.3895 | -101.0552
5.3932 | -101.0062
5.3969 | -100.9514
5.4007 | -100.8947
5.4044 | -100.8411
5.4081 | -100.7875
5.4119 | -100.7509
5.4156 | -100.7369
5.4193 | -100.7425
5.423 | -100.7662
5.4268 | -100.8047
5.4305 | -100.8583
5.4342 | -100.9244
5.4379 | -101.0023
5.4417 | -101.0795
5.4454 | -101.1624
5.4491 | -101.2414
5.4529 | -101.3091
5.4566 | -101.3688
5.4603 | -101.4242
5.464 | -101.4574
5.4678 | -101.4717
5.4715 | -101.4761
5.4752 | -101.4632
5.4789 | -101.4389
5.4827 | -101.4097
5.4864 | -101.3868
5.4901 | -101.3691
5.4939 | -101.3642
5.4976 | -101.3743
5.5013 | -101.4042
5.505 | -101.4461
5.5088 | -101.49
5.5125 | -101.5389
5.5162 | -101.5868
5.5199 | -101.6216
5.5237 | -101.6408
5.5274 | -101.642
5.5311 | -101.6231
5.5348 | -101.5978
5.5386 | -101.5672
5.5423 | -101.5379
5.546 | -101.5082
5.5498 | -101.4839
5.5535 | -101.4535
5.5572 | -101.4255
5.5609 | -101.4057
5.5647 | -101.3838
5.5684 | -101.3573
5.5721 | -101.3291
5.5758 | -101.3145
5.5796 | -101.3094
5.5833 | -101.306
5.587 | -101.3002
5.5908 | -101.2986
5.5945 | -101.3025
5.5982 | -101.2975
5.6019 | -101.2907
5.6057 | -101.2801
5.6094 | -101.2674
5.6131 | -101.2578
5.6168 | -101.2547
5.6206 | -101.2508
5.6243 | -101.251
5.628 | -101.2607
5.6318 | -101.2888
5.6355 | -101.336
5.6392 | -101.4038
5.6429 | -101.4767
5.6467 | -101.5615
5.6504 | -101.6473
5.6541 | -101.7315
5.6578 | -101.8104
5.6616 | -101.8624
5.6653 | -101.8953
5.669 | -101.9148
5.6728 | -101.9154
5.6765 | -101.8936
5.6802 | -101.854
5.6839 | -101.8058
5.6877 | -101.7502
5.6914 | -101.6983
5.6951 | -101.6566
5.6988 | -101.6226
5.7026 | -101.5948
5.7063 | -101.5776
5.71 | -101.569
5.7138 | -101.5662
5.7175 | -101.5634
5.7212 | -101.5612
5.7249 | -101.5523
5.7287 | -101.5392
5.7324 | -101.5183
5.7361 | -101.5023
5.7398 | -101.4835
5.7436 | -101.4618
5.7473 | -101.4493
5.751 | -101.4378
5.7548 | -101.4328
5.7585 | -101.4354
5.7622 | -101.4427
5.7659 | -101.4452
5.7697 | -101.4437
5.7734 | -101.4431
5.7771 | -101.4301
5.7808 | -101.41
5.7846 | -101.3902
5.7883 | -101.3708
5.792 | -101.3588
5.7958 | -101.3454
5.7995 | -101.3354
5.8032 | -101.3383
5.8069 | -101.3496
5.8107 | -101.371
5.8144 | -101.3949
5.8181 | -101.426
5.8218 | -101.4619
5.8256 | -101.5059
5.8293 | -101.5502
5.833 | -101.5937
5.8368 | -101.6414
5.8405 | -101.6857
5.8442 | -101.725
5.8479 | -101.7617
5.8517 | -101.7944
5.8554 | -101.8184
5.8591 | -101.8351
5.8628 | -101.8531
5.8666 | -101.8643
5.8703 | -101.8678
5.874 | -101.8704
5.8777 | -101.8661
5.8815 | -101.8637
5.8852 | -101.863
5.8889 | -101.8593
5.8927 | -101.8504
5.8964 | -101.8415
5.9001 | -101.8313
5.9038 | -101.8059
5.9076 | -101.7733
5.9113 | -101.7374
5.915 | -101.7007
5.9187 | -101.6672
5.9225 | -101.6374
5.9262 | -101.607
5.9299 | -101.575
5.9337 | -101.5456
5.9374 | -101.511
5.9411 | -101.4733
5.9448 | -101.4422
5.9486 | -101.412
5.9523 | -101.3865
5.956 | -101.3713
5.9597 | -101.3669
5.9635 | -101.3716
5.9672 | -101.3891
5.9709 | -101.4215
5.9747 | -101.462
5.9784 | -101.5107
5.9821 | -101.5702
5.9858 | -101.6316
5.9896 | -101.6982
5.9933 | -101.7655
5.997 | -101.8259
6.0007 | -101.8761
6.0045 | -101.9143
6.0082 | -101.9451
6.0119 | -101.9612
6.0157 | -101.9688
6.0194 | -101.9762
6.0231 | -101.9879
6.0268 | -102.0044
6.0306 | -102.0343
6.0343 | -102.0801
6.038 | -102.1252
6.0417 | -102.1734
6.0455 | -102.2292
6.0492 | -102.2823
6.0529 | -102.3389
6.0567 | -102.3896
6.0604 | -102.424
6.0641 | -102.4409
6.0678 | -102.4423
6.0716 | -102.4324
6.0753 | -102.4121
6.079 | -102.3835
6.0827 | -102.3508
6.0865 | -102.3195
6.0902 | -102.2945
6.0939 | -102.2754
6.0977 | -102.2628
6.1014 | -102.2572
6.1051 | -102.2617
6.1088 | -102.2761
6.1126 | -102.2904
6.1163 | -102.3015
6.12 | -102.3168
6.1237 | -102.3272
6.1275 | -102.3327
6.1312 | -102.3393
6.1349 | -102.349
6.1387 | -102.3494
6.1424 | -102.3507
6.1461 | -102.3618
6.1498 | -102.3787
6.1536 | -102.4068
6.1573 | -102.4352
6.161 | -102.4639
6.1647 | -102.4963
6.1685 | -102.526
6.1722 | -102.5567
6.1759 | -102.5837
6.1796 | -102.61
6.1834 | -102.6317
6.1871 | -102.6512
6.1908 | -102.6723
6.1946 | -102.6851
6.1983 | -102.6847
6.202 | -102.6749
6.2057 | -102.6562
6.2095 | -102.6295
6.2132 | -102.5913
6.2169 | -102.5481
6.2206 | -102.4886
6.2244 | -102.417
6.2281 | -102.3359
6.2318 | -102.2588
6.2356 | -102.1827
6.2393 | -102.1036
6.243 | -102.0332
6.2467 | -101.9776
6.2505 | -101.9434
6.2542 | -101.9153
6.2579 | -101.8965
6.2616 | -101.8918
6.2654 | -101.891
6.2691 | -101.8971
6.2728 | -101.9092
6.2766 | -101.9212
6.2803 | -101.9327
6.284 | -101.9395
6.2877 | -101.9498
6.2915 | -101.9674
6.2952 | -101.9818
6.2989 | -102.0017
6.3026 | -102.0252
6.3064 | -102.0446
6.3101 | -102.0648
6.3138 | -102.0907
6.3176 | -102.1188
6.3213 | -102.1306
6.325 | -102.1247
6.3287 | -102.1083
6.3325 | -102.0821
6.3362 | -102.0496
6.3399 | -102.0055
6.3436 | -101.9482
6.3474 | -101.8996
6.3511 | -101.8517
6.3548 | -101.8072
6.3586 | -101.779
6.3623 | -101.7627
6.366 | -101.7565
6.3697 | -101.7682
6.3735 | -101.7918
6.3772 | -101.8263
6.3809 | -101.8691
6.3846 | -101.9195
6.3884 | -101.9825
6.3921 | -102.044
6.3958 | -102.1124
6.3996 | -102.1862
6.4033 | -102.2514
6.407 | -102.3089
6.4107 | -102.3513
6.4145 | -102.3879
6.4182 | -102.4063
6.4219 | -102.4057
6.4256 | -102.3919
6.4294 | -102.3618
6.4331 | -102.3209
6.4368 | -102.2794
6.4406 | -102.2321
6.4443 | -102.1829
6.448 | -102.1362
6.4517 | -102.1237
6.4555 | -102.0997
6.4592 | -102.084
6.4629 | -102.0845
6.4666 | -102.0939
6.4704 | -102.1132
6.4741 | -102.1401
6.4778 | -102.1742
6.4816 | -102.2138
6.4853 | -102.2618
6.489 | -102.3151
6.4927 | -102.3672
6.4965 | -102.4246
6.5002 | -102.4902
6.5039 | -102.5612
6.5076 | -102.6323
6.5114 | -102.7059
6.5151 | -102.7779
6.5188 | -102.8462
6.5225 | -102.9148
6.5263 | -102.9798
6.53 | -103.0419
6.5337 | -103.0932
6.5375 | -103.1378
6.5412 | -103.1767
6.5449 | -103.2075
6.5486 | -103.2237
6.5524 | -103.2233
6.5561 | -103.2162
6.5598 | -103.2025
6.5635 | -103.1732
6.5673 | -103.1317
6.571 | -103.0817
6.5747 | -103.0273
6.5785 | -102.9697
6.5822 | -102.9133
6.5859 | -102.8516
6.5896 | -102.7905
6.5934 | -102.7256
6.5971 | -102.6622
6.6008 | -102.5987
6.6045 | -102.5354
6.6083 | -102.4801
6.612 | -102.4258
6.6157 | -102.3826
6.6195 | -102.3521
6.6232 | -102.3321
6.6269 | -102.327
6.6306 | -102.3309
6.6344 | -102.3502
6.6381 | -102.3677
6.6418 | -102.3909
6.6455 | -102.4102
6.6493 | -102.4194
6.653 | -102.4178
6.6567 | -102.4089
6.6605 | -102.3872
6.6642 | -102.3614
6.6679 | -102.3394
6.6716 | -102.3178
6.6754 | -102.3004
6.6791 | -102.2804
6.6828 | -102.261
6.6865 | -102.2572
6.6903 | -102.2675
6.694 | -102.2859
6.6977 | -102.3066
6.7015 | -102.34
6.7052 | -102.3828
6.7089 | -102.4277
6.7126 | -102.4723
6.7164 | -102.5072
6.7201 | -102.5362
6.7238 | -102.557
6.7275 | -102.569
6.7313 | -102.5681
6.735 | -102.5545
6.7387 | -102.5389
6.7425 | -102.5119
6.7462 | -102.4878
6.7499 | -102.4739
6.7536 | -102.4663
6.7574 | -102.4563
6.7611 | -102.45
6.7648 | -102.4515
6.7685 | -102.4551
6.7723 | -102.4648
6.776 | -102.4781
6.7797 | -102.4986
6.7835 | -102.5199
6.7872 | -102.534
6.7909 | -102.544
6.7946 | -102.5559
6.7984 | -102.5684
6.8021 | -102.5746
6.8058 | -102.5836
6.8095 | -102.6
6.8133 | -102.6238
6.817 | -102.6554
6.8207 | -102.6857
6.8244 | -102.7138
6.8282 | -102.7395
6.8319 | -102.7613
6.8356 | -102.7742
6.8394 | -102.7796
6.8431 | -102.7771
6.8468 | -102.7648
6.8505 | -102.7407
6.8543 | -102.7095
6.858 | -102.677
6.8617 | -102.6444
6.8654 | -102.6038
6.8692 | -102.5614
6.8729 | -102.5254
6.8766 | -102.4911
6.8804 | -102.4585
6.8841 | -102.4379
6.8878 | -102.4267
6.8915 | -102.4282
6.8953 | -102.4462
6.899 | -102.4794
6.9027 | -102.5287
6.9064 | -102.5895
6.9102 | -102.6631
6.9139 | -102.7407
6.9176 | -102.8167
6.9214 | -102.8933
6.9251 | -102.9634
6.9288 | -103.0217
6.9325 | -103.071
6.9363 | -103.0997
6.94 | -103.103
6.9437 | -103.0824
6.9474 | -103.0383
6.9512 | -102.9482
6.9549 | -102.8477
6.9586 | -102.7425
6.9624 | -102.625
6.9661 | -102.5007
6.9698 | -102.3704
6.9735 | -102.2366
6.9773 | -102.1149
6.981 | -102.0073
6.9847 | -101.9141
6.9884 | -101.9049
6.9922 | -101.8695
6.9959 | -101.8573
6.9996 | -101.8718
7.0034 | -101.9129
7.0071 | -101.9731
7.0108 | -102.0502
7.0145 | -102.1401
7.0183 | -102.2425
7.022 | -102.3507
7.0257 | -102.4551
7.0294 | -102.5547
7.0332 | -102.6386
7.0369 | -102.7047
7.0406 | -102.7618
7.0444 | -102.8105
7.0481 | -102.8491
7.0518 | -102.8726
7.0555 | -102.8959
7.0593 | -102.8923
7.063 | -102.874
7.0667 | -102.8518
7.0704 | -102.8272
7.0742 | -102.8025
7.0779 | -102.7789
7.0816 | -102.7592
7.0854 | -102.7447
7.0891 | -102.7393
7.0928 | -102.742
7.0965 | -102.7572
7.1003 | -102.7828
7.104 | -102.8177
7.1077 | -102.8542
7.1114 | -102.8969
7.1152 | -102.9439
7.1189 | -102.9919
7.1226 | -103.0301
7.1264 | -103.0656
7.1301 | -103.0888
7.1338 | -103.1044
7.1375 | -103.1111
7.1413 | -103.1135
7.145 | -103.1101
7.1487 | -103.1001
7.1524 | -103.0827
7.1562 | -103.0654
7.1599 | -103.0579
7.1636 | -103.0542
7.1673 | -103.0509
7.1711 | -103.0479
7.1748 | -103.0534
7.1785 | -103.0603
7.1823 | -103.0658
7.186 | -103.0723
7.1897 | -103.0733
7.1934 | -103.074
7.1972 | -103.0719
7.2009 | -103.0679
7.2046 | -103.0576
7.2083 | -103.0453
7.2121 | -103.0331
7.2158 | -103.0218
7.2195 | -103.0088
7.2233 | -102.9997
7.227 | -102.992
7.2307 | -102.989
7.2344 | -102.9906
7.2382 | -102.9944
7.2419 | -102.9997
7.2456 | -103.0167
7.2493 | -103.0321
7.2531 | -103.0488
7.2568 | -103.0662
7.2605 | -103.0804
7.2643 | -103.0928
7.268 | -103.1023
7.2717 | -103.1128
7.2754 | -103.123
7.2792 | -103.1295
7.2829 | -103.1354
7.2866 | -103.1483
7.2903 | -103.1686
7.2941 | -103.1897
7.2978 | -103.2143
7.3015 | -103.246
7.3053 | -103.2776
7.309 | -103.3151
7.3127 | -103.3592
7.3164 | -103.4016
7.3202 | -103.4399
7.3239 | -103.4737
7.3276 | -103.5009
7.3313 | -103.5142
7.3351 | -103.5123
7.3388 | -103.5052
7.3425 | -103.4766
7.3463 | -103.4428
7.35 | -103.391
7.3537 | -103.3285
7.3574 | -103.2745
7.3612 | -103.2128
7.3649 | -103.1553
7.3686 | -103.1096
7.3723 | -103.0683
7.3761 | -103.0316
7.3798 | -103.0069
7.3835 | -102.9909
7.3873 | -102.9793
7.391 | -102.9789
7.3947 | -102.9889
7.3984 | -103.0035
7.4022 | -103.019
7.4059 | -103.0358
7.4096 | -103.0559
7.4133 | -103.0798
7.4171 | -103.1042
7.4208 | -103.1258
7.4245 | -103.143
7.4283 | -103.1523
7.432 | -103.1525
7.4357 | -103.1481
7.4394 | -103.1411
7.4432 | -103.1281
7.4469 | -103.1075
7.4506 | -103.0888
7.4543 | -103.0666
7.4581 | -103.0455
7.4618 | -103.0198
7.4655 | -102.9886
7.4693 | -102.9669
7.473 | -102.9469
7.4767 | -102.9292
7.4804 | -102.9182
7.4842 | -102.9161
7.4879 | -102.9182
7.4916 | -102.9286
7.4953 | -102.9497
7.4991 | -102.9695
7.5028 | -102.989
7.5065 | -103.0125
7.5103 | -103.0442
7.514 | -103.0829
7.5177 | -103.1171
7.5214 | -103.1531
7.5252 | -103.1956
7.5289 | -103.2274
7.5326 | -103.261
7.5363 | -103.2953
7.5401 | -103.3307
7.5438 | -103.3595
7.5475 | -103.3798
7.5512 | -103.391
7.555 | -103.3892
7.5587 | -103.379
7.5624 | -103.3621
7.5662 | -103.3393
7.5699 | -103.3089
7.5736 | -103.2745
7.5773 | -103.2413
7.5811 | -103.2008
7.5848 | -103.1621
7.5885 | -103.1244
7.5922 | -103.0854
7.596 | -103.0478
7.5997 | -103.0147
7.6034 | -102.9856
7.6072 | -102.9623
7.6109 | -102.949
7.6146 | -102.9409
7.6183 | -102.9393
7.6221 | -102.9446
7.6258 | -102.956
7.6295 | -102.9758
7.6332 | -103.0077
7.637 | -103.05
7.6407 | -103.0991
7.6444 | -103.1542
7.6482 | -103.2154
7.6519 | -103.2728
7.6556 | -103.3259
7.6593 | -103.3818
7.6631 | -103.4383
7.6668 | -103.488
7.6705 | -103.5311
7.6742 | -103.562
7.678 | -103.5745
7.6817 | -103.571
7.6854 | -103.5528
7.6892 | -103.521
7.6929 | -103.4812
7.6966 | -103.4353
7.7003 | -103.3796
7.7041 | -103.3244
7.7078 | -103.2712
7.7115 | -103.2233
7.7152 | -103.186
7.719 | -103.1563
7.7227 | -103.1316
7.7264 | -103.1199
7.7302 | -103.126
7.7339 | -103.1444
7.7376 | -103.1717
7.7413 | -103.2086
7.7451 | -103.2506
7.7488 | -103.2889
7.7525 | -103.3245
7.7562 | -103.3641
7.76 | -103.3981
7.7637 | -103.4223
7.7674 | -103.4436
7.7712 | -103.4526
7.7749 | -103.4432
7.7786 | -103.4234
7.7823 | -103.4036
7.7861 | -103.3852
7.7898 | -103.3629
7.7935 | -103.348
7.7972 | -103.3374
7.801 | -103.3333
7.8047 | -103.3343
7.8084 | -103.3404
7.8122 | -103.3573
7.8159 | -103.3852
7.8196 | -103.4177
7.8233 | -103.4578
7.8271 | -103.4969
7.8308 | -103.5308
7.8345 | -103.5713
7.8382 | -103.6187
7.842 | -103.6564
7.8457 | -103.6864
7.8494 | -103.7074
7.8531 | -103.7194
7.8569 | -103.7235
7.8606 | -103.7212
7.8643 | -103.7185
7.8681 | -103.7124
7.8718 | -103.7001
7.8755 | -103.6862
7.8792 | -103.669
7.883 | -103.6535
7.8867 | -103.6423
7.8904 | -103.6323
7.8941 | -103.6206
7.8979 | -103.6085
7.9016 | -103.5944
7.9053 | -103.5829
7.9091 | -103.5723
7.9128 | -103.5621
7.9165 | -103.5603
7.9202 | -103.56
7.924 | -103.5639
7.9277 | -103.5736
7.9314 | -103.582
7.9351 | -103.5924
7.9389 | -103.6042
7.9426 | -103.6172
7.9463 | -103.6295
7.9501 | -103.6297
7.9538 | -103.6198
7.9575 | -103.6052
7.9612 | -103.5852
7.965 | -103.5613
7.9687 | -103.5375
7.9724 | -103.508
7.9761 | -103.4805
7.9799 | -103.4524
7.9836 | -103.4297
7.9873 | -103.4152
7.9911 | -103.4123
7.9948 | -103.4185
7.9985 | -103.4314
8.0022 | -103.4504
8.006 | -103.4757
8.0097 | -103.508
8.0134 | -103.5444
8.0171 | -103.5805
8.0209 | -103.6107
8.0246 | -103.6363
8.0283 | -103.6551
8.0321 | -103.6561
8.0358 | -103.6518
8.0395 | -103.6411
8.0432 | -103.6225
8.047 | -103.5995
8.0507 | -103.5715
8.0544 | -103.5414
8.0581 | -103.5129
8.0619 | -103.4837
8.0656 | -103.4817
8.0693 | -103.4695
8.0731 | -103.4634
8.0768 | -103.464
8.0805 | -103.4712
8.0842 | -103.4829
8.088 | -103.4919
8.0917 | -103.5087
8.0954 | -103.5297
8.0991 | -103.5489
8.1029 | -103.5679
8.1066 | -103.586
8.1103 | -103.6028
8.1141 | -103.6173
8.1178 | -103.6283
8.1215 | -103.6345
8.1252 | -103.6406
8.129 | -103.6426
8.1327 | -103.6428
8.1364 | -103.6437
8.1401 | -103.6464
8.1439 | -103.6459
8.1476 | -103.6526
8.1513 | -103.6653
8.155 | -103.6783
8.1588 | -103.6952
8.1625 | -103.7188
8.1662 | -103.7546
8.17 | -103.7965
8.1737 | -103.8457
8.1774 | -103.9003
8.1811 | -103.9576
8.1849 | -104.016
8.1886 | -104.072
8.1923 | -104.1285
8.196 | -104.1782
8.1998 | -104.2182
8.2035 | -104.2514
8.2072 | -104.2744
8.211 | -104.2849
8.2147 | -104.2885
8.2184 | -104.2767
8.2221 | -104.2555
8.2259 | -104.2267
8.2296 | -104.1901
8.2333 | -104.1443
8.237 | -104.0956
8.2408 | -104.0493
8.2445 | -104.0017
8.2482 | -103.9548
8.252 | -103.9128
8.2557 | -103.8747
8.2594 | -103.8449
8.2631 | -103.8214
8.2669 | -103.803
8.2706 | -103.7891
8.2743 | -103.783
8.278 | -103.7832
8.2818 | -103.7825
8.2855 | -103.7904
8.2892 | -103.7997
8.293 | -103.8138
8.2967 | -103.8266
8.3004 | -103.8401
8.3041 | -103.8484
8.3079 | -103.8502
8.3116 | -103.8475
8.3153 | -103.8409
8.319 | -103.835
8.3228 | -103.8241
8.3265 | -103.8065
8.3302 | -103.7895
8.334 | -103.7737
8.3377 | -103.7491
8.3414 | -103.73
8.3451 | -103.719
8.3489 | -103.713
8.3526 | -103.7112
8.3563 | -103.7166
8.36 | -103.7215
8.3638 | -103.7311
8.3675 | -103.7514
8.3712 | -103.7751
8.375 | -103.7988
8.3787 | -103.8279
8.3824 | -103.8609
8.3861 | -103.8945
8.3899 | -103.9333
8.3936 | -103.97
8.3973 | -104.0059
8.401 | -104.0438
8.4048 | -104.0868
8.4085 | -104.124
8.4122 | -104.1553
8.416 | -104.1849
8.4197 | -104.2128
8.4234 | -104.239
8.4271 | -104.2675
8.4309 | -104.2907
8.4346 | -104.3058
8.4383 | -104.3103
8.442 | -104.305
8.4458 | -104.3007
8.4495 | -104.2966
8.4532 | -104.2848
8.457 | -104.2717
8.4607 | -104.2643
8.4644 | -104.2528
8.4681 | -104.2364
8.4719 | -104.2145
8.4756 | -104.1981
8.4793 | -104.1821
8.483 | -104.1715
8.4868 | -104.1612
8.4905 | -104.1527
8.4942 | -104.1479
8.4979 | -104.1471
8.5017 | -104.151
8.5054 | -104.1588
8.5091 | -104.1711
8.5129 | -104.1794
8.5166 | -104.1871
8.5203 | -104.1938
8.524 | -104.1999
8.5278 | -104.2015
8.5315 | -104.2026
8.5352 | -104.2051
8.5389 | -104.1974
8.5427 | -104.19
8.5464 | -104.1778
8.5501 | -104.1602
8.5539 | -104.1408
8.5576 | -104.1244
8.5613 | -104.1113
8.565 | -104.0936
8.5688 | -104.0712
8.5725 | -104.049
8.5762 | -104.0258
8.5799 | -104.0029
8.5837 | -103.9797
8.5874 | -103.9612
8.5911 | -103.9491
8.5949 | -103.9451
8.5986 | -103.9412
8.6023 | -103.9578
8.606 | -103.9658
8.6098 | -103.9765
8.6135 | -103.988
8.6172 | -104.0026
8.6209 | -104.022
8.6247 | -104.042
8.6284 | -104.0575
8.6321 | -104.0672
8.6359 | -104.0775
8.6396 | -104.0912
8.6433 | -104.0985
8.647 | -104.1024
8.6508 | -104.1026
8.6545 | -104.1016
8.6582 | -104.098
8.6619 | -104.0928
8.6657 | -104.0969
8.6694 | -104.0936
8.6731 | -104.0889
8.6769 | -104.0835
8.6806 | -104.075
8.6843 | -104.064
8.688 | -104.0534
8.6918 | -104.0442
8.6955 | -104.0315
8.6992 | -104.01
8.7029 | -103.9879
8.7067 | -103.9619
8.7104 | -103.9315
8.7141 | -103.8941
8.7179 | -103.8592
8.7216 | -103.8238
8.7253 | -103.7879
8.729 | -103.7529
8.7328 | -103.7224
8.7365 | -103.6949
8.7402 | -103.673
8.7439 | -103.6529
8.7477 | -103.6378
8.7514 | -103.6284
8.7551 | -103.6229
8.7589 | -103.6278
8.7626 | -103.6371
8.7663 | -103.6446
8.77 | -103.652
8.7738 | -103.6701
8.7775 | -103.6927
8.7812 | -103.7153
8.7849 | -103.7424
8.7887 | -103.7716
8.7924 | -103.8069
8.7961 | -103.8365
8.7999 | -103.8702
8.8036 | -103.9077
8.8073 | -103.9443
8.811 | -103.9806
8.8148 | -104.0137
8.8185 | -104.048
8.8222 | -104.0813
8.8259 | -104.1109
8.8297 | -104.1446
8.8334 | -104.1801
8.8371 | -104.21
8.8408 | -104.2392
8.8446 | -104.2694
8.8483 | -104.2906
8.852 | -104.3109
8.8558 | -104.3291
8.8595 | -104.3422
8.8632 | -104.3552
8.8669 | -104.3676
8.8707 | -104.3759
8.8744 | -104.3883
8.8781 | -104.4007
8.8818 | -104.412
8.8856 | -104.4295
8.8893 | -104.4426
8.893 | -104.457
8.8968 | -104.474
8.9005 | -104.4907
8.9042 | -104.4989
8.9079 | -104.5023
8.9117 | -104.5063
8.9154 | -104.5026
8.9191 | -104.496
8.9228 | -104.4847
8.9266 | -104.4701
8.9303 | -104.4447
8.934 | -104.4128
8.9378 | -104.374
8.9415 | -104.3426
8.9452 | -104.3096
8.9489 | -104.2751
8.9527 | -104.2443
8.9564 | -104.2196
8.9601 | -104.1941
8.9638 | -104.1664
8.9676 | -104.1429
8.9713 | -104.1214
8.975 | -104.1043
8.9788 | -104.0901
8.9825 | -104.0772
8.9862 | -104.0681
8.9899 | -104.0628
8.9937 | -104.0593
8.9974 | -104.0605
9.0011 | -104.0638
9.0048 | -104.0698
9.0086 | -104.0794
9.0123 | -104.0946
9.016 | -104.1112
9.0198 | -104.1266
9.0235 | -104.1466
9.0272 | -104.1679
9.0309 | -104.1924
9.0347 | -104.2152
9.0384 | -104.2382
9.0421 | -104.2599
9.0458 | -104.2791
9.0496 | -104.2936
9.0533 | -104.3017
9.057 | -104.3159
9.0608 | -104.3372
9.0645 | -104.3512
9.0682 | -104.3599
9.0719 | -104.3539
9.0757 | -104.3515
9.0794 | -104.3461
9.0831 | -104.3355
9.0868 | -104.327
9.0906 | -104.3205
9.0943 | -104.3171
9.098 | -104.312
9.1018 | -104.3074
9.1055 | -104.3009
9.1092 | -104.2954
9.1129 | -104.2923
9.1167 | -104.2956
9.1204 | -104.2994
9.1241 | -104.3032
9.1278 | -104.3064
9.1316 | -104.316
9.1353 | -104.3252
9.139 | -104.3362
9.1428 | -104.3399
9.1465 | -104.3415
9.1502 | -104.3462
9.1539 | -104.3536
9.1577 | -104.3593
9.1614 | -104.366
9.1651 | -104.3697
9.1688 | -104.3785
9.1726 | -104.389
9.1763 | -104.3999
9.18 | -104.4154
9.1837 | -104.43
9.1875 | -104.4464
9.1912 | -104.4577
9.1949 | -104.4647
9.1987 | -104.4757
9.2024 | -104.4898
9.2061 | -104.5057
9.2098 | -104.5191
9.2136 | -104.5312
9.2173 | -104.5464
9.221 | -104.5586
9.2247 | -104.5691
9.2285 | -104.5785
9.2322 | -104.5883
9.2359 | -104.5966
9.2397 | -104.6115
9.2434 | -104.6252
9.2471 | -104.6349
9.2508 | -104.6392
9.2546 | -104.6442
9.2583 | -104.6493
9.262 | -104.6571
9.2657 | -104.6574
9.2695 | -104.6577
9.2732 | -104.6515
9.2769 | -104.6439
9.2807 | -104.6445
9.2844 | -104.6425
9.2881 | -104.638
9.2918 | -104.633
9.2956 | -104.6271
9.2993 | -104.6156
9.303 | -104.604
9.3067 | -104.5915
9.3105 | -104.5777
9.3142 | -104.5686
9.3179 | -104.5613
9.3217 | -104.5552
9.3254 | -104.5502
9.3291 | -104.5446
9.3328 | -104.5396
9.3366 | -104.5326
9.3403 | -104.5299
9.344 | -104.5294
9.3477 | -104.5306
9.3515 | -104.5385
9.3552 | -104.5469
9.3589 | -104.5563
9.3627 | -104.5696
9.3664 | -104.5875
9.3701 | -104.6058
9.3738 | -104.6242
9.3776 | -104.6347
9.3813 | -104.6475
9.385 | -104.6606
9.3887 | -104.6728
9.3925 | -104.6813
9.3962 | -104.6836
9.3999 | -104.6833
9.4037 | -104.6836
9.4074 | -104.6785
9.4111 | -104.6734
9.4148 | -104.6629
9.4186 | -104.6427
9.4223 | -104.6273
9.426 | -104.6104
9.4297 | -104.5896
9.4335 | -104.569
9.4372 | -104.5525
9.4409 | -104.5361
9.4447 | -104.5199
9.4484 | -104.5014
9.4521 | -104.4817
9.4558 | -104.4626
9.4596 | -104.4458
9.4633 | -104.4307
9.467 | -104.4167
9.4707 | -104.4006
9.4745 | -104.3879
9.4782 | -104.3741
9.4819 | -104.3599
9.4856 | -104.3499
9.4894 | -104.342
9.4931 | -104.3386
9.4968 | -104.343
9.5006 | -104.3521
9.5043 | -104.3618
9.508 | -104.3782
9.5117 | -104.3993
9.5155 | -104.4269
9.5192 | -104.4526
9.5229 | -104.4751
9.5266 | -104.4992
9.5304 | -104.5229
9.5341 | -104.5456
9.5378 | -104.5638
9.5416 | -104.5783
9.5453 | -104.5926
9.549 | -104.6036
9.5527 | -104.6171
9.5565 | -104.6252
9.5602 | -104.6302
9.5639 | -104.637
9.5676 | -104.6462
9.5714 | -104.6547
9.5751 | -104.668
9.5788 | -104.6849
9.5826 | -104.7021
9.5863 | -104.7245
9.59 | -104.7469
9.5937 | -104.765
9.5975 | -104.7849
9.6012 | -104.8035
9.6049 | -104.8167
9.6086 | -104.8302
9.6124 | -104.8378
9.6161 | -104.8405
9.6198 | -104.8397
9.6236 | -104.8311
9.6273 | -104.8196
9.631 | -104.7992
9.6347 | -104.7763
9.6385 | -104.7483
9.6422 | -104.7212
9.6459 | -104.6867
9.6496 | -104.6478
9.6534 | -104.6091
9.6571 | -104.5707
9.6608 | -104.536
9.6646 | -104.5002
9.6683 | -104.4653
9.672 | -104.436
9.6757 | -104.4237
9.6795 | -104.4064
9.6832 | -104.3929
9.6869 | -104.3783
9.6906 | -104.3697
9.6944 | -104.3675
9.6981 | -104.3711
9.7018 | -104.3813
9.7056 | -104.396
9.7093 | -104.414
9.713 | -104.4354
9.7167 | -104.4552
9.7205 | -104.4782
9.7242 | -104.5035
9.7279 | -104.525
9.7316 | -104.5471
9.7354 | -104.5733
9.7391 | -104.5994
9.7428 | -104.621
9.7466 | -104.6412
9.7503 | -104.6584
9.754 | -104.6728
9.7577 | -104.68
9.7615 | -104.6882
9.7652 | -104.6945
9.7689 | -104.6935
9.7726 | -104.6924
9.7764 | -104.6853
9.7801 | -104.6751
9.7838 | -104.6635
9.7876 | -104.6577
9.7913 | -104.6498
9.795 | -104.6388
9.7987 | -104.6313
9.8025 | -104.6261
9.8062 | -104.6193
9.8099 | -104.6141
9.8136 | -104.6083
9.8174 | -104.6031
9.8211 | -104.5995
9.8248 | -104.5979
9.8285 | -104.6
9.8323 | -104.6035
9.836 | -104.6065
9.8397 | -104.6078
9.8435 | -104.6095
9.8472 | -104.6068
9.8509 | -104.6051
9.8546 | -104.6025
9.8584 | -104.5988
9.8621 | -104.5964
9.8658 | -104.5943
9.8695 | -104.5921
9.8733 | -104.5922
9.877 | -104.59
9.8807 | -104.5811
9.8845 | -104.5705
9.8882 | -104.5595
9.8919 | -104.5506
9.8956 | -104.5411
9.8994 | -104.532
9.9031 | -104.5241
9.9068 | -104.5098
9.9105 | -104.4934
9.9143 | -104.4803
9.918 | -104.4665
9.9217 | -104.4535
9.9255 | -104.4422
9.9292 | -104.4288
9.9329 | -104.4143
9.9366 | -104.4029
9.9404 | -104.3963
9.9441 | -104.3912
9.9478 | -104.3862
9.9515 | -104.384
9.9553 | -104.3899
9.959 | -104.4002
9.9627 | -104.4137
9.9665 | -104.427
9.9702 | -104.441
9.9739 | -104.4574
9.9776 | -104.4789
9.9814 | -104.5004
9.9851 | -104.5192
9.9888 | -104.5345
9.9925 | -104.5495
9.9963 | -104.5623
10 | -104.5749
10.0037 | -104.5837
10.0075 | -104.5919
10.0112 | -104.5981
10.0149 | -104.6049
10.0186 | -104.6113
10.0224 | -104.6192
10.0261 | -104.6301
10.0298 | -104.6399
10.0335 | -104.6524
10.0373 | -104.6629
10.041 | -104.6754
10.0447 | -104.6894
10.0485 | -104.7036
10.0522 | -104.7213
10.0559 | -104.7401
10.0596 | -104.7536
10.0634 | -104.7634
10.0671 | -104.7708
10.0708 | -104.779
10.0745 | -104.7853
10.0783 | -104.7898
10.082 | -104.7899
10.0857 | -104.7866
10.0895 | -104.7768
10.0932 | -104.7597
10.0969 | -104.738
10.1006 | -104.7097
10.1044 | -104.6771
10.1081 | -104.6446
10.1118 | -104.6131
10.1155 | -104.5793
10.1193 | -104.5455
10.123 | -104.5137
10.1267 | -104.4827
10.1305 | -104.4598
10.1342 | -104.4463
10.1379 | -104.4341
10.1416 | -104.4295
10.1454 | -104.429
10.1491 | -104.434
10.1528 | -104.4427
10.1565 | -104.46
10.1603 | -104.4842
10.164 | -104.514
10.1677 | -104.5444
10.1715 | -104.5803
10.1752 | -104.6177
10.1789 | -104.6553
10.1826 | -104.6921
10.1864 | -104.7272
10.1901 | -104.7616
10.1938 | -104.7957
10.1975 | -104.8225
10.2013 | -104.8461
10.205 | -104.8688
10.2087 | -104.8872
10.2124 | -104.8977
10.2162 | -104.8903
10.2199 | -104.8878
10.2236 | -104.8825
10.2274 | -104.8761
10.2311 | -104.8693
10.2348 | -104.8586
10.2385 | -104.8499
10.2423 | -104.8425
10.246 | -104.8284
10.2497 | -104.8126
10.2534 | -104.7987
10.2572 | -104.7822
10.2609 | -104.7651
10.2646 | -104.7512
10.2684 | -104.7361
10.2721 | -104.7175
10.2758 | -104.7041
10.2795 | -104.6902
10.2833 | -104.6765
10.287 | -104.6673
10.2907 | -104.6585
10.2944 | -104.6542
10.2982 | -104.6536
10.3019 | -104.6557
10.3056 | -104.6587
10.3094 | -104.6649
10.3131 | -104.675
10.3168 | -104.6888
10.3205 | -104.7041
10.3243 | -104.7214
10.328 | -104.7431
10.3317 | -104.7648
10.3354 | -104.7862
10.3392 | -104.8054
10.3429 | -104.8271
10.3466 | -104.8528
10.3504 | -104.875
10.3541 | -104.8955
10.3578 | -104.9157
10.3615 | -104.933
10.3653 | -104.9508
10.369 | -104.9642
10.3727 | -104.9764
10.3764 | -104.983
10.3802 | -104.9854
10.3839 | -104.9879
10.3876 | -104.9867
10.3914 | -104.9848
10.3951 | -104.9809
10.3988 | -104.9802
10.4025 | -104.9715
10.4063 | -104.9551
10.41 | -104.9421
10.4137 | -104.9312
10.4174 | -104.9186
10.4212 | -104.9069
10.4249 | -104.8977
10.4286 | -104.8835
10.4324 | -104.8723
10.4361 | -104.8638
10.4398 | -104.8549
10.4435 | -104.8501
10.4473 | -104.8488
10.451 | -104.848
10.4547 | -104.8476
10.4584 | -104.8517
10.4622 | -104.8539
10.4659 | -104.8551
10.4696 | -104.855
10.4734 | -104.8546
10.4771 | -104.8501
10.4808 | -104.8429
10.4845 | -104.8376
10.4883 | -104.8287
10.492 | -104.8195
10.4957 | -104.8044
10.4994 | -104.787
10.5032 | -104.7707
10.5069 | -104.7512
10.5106 | -104.7308
10.5144 | -104.7112
10.5181 | -104.6942
10.5218 | -104.6815
10.5255 | -104.6703
10.5293 | -104.6619
10.533 | -104.6543
10.5367 | -104.6451
10.5404 | -104.6384
10.5442 | -104.6339
10.5479 | -104.6341
10.5516 | -104.635
10.5553 | -104.6398
10.5591 | -104.647
10.5628 | -104.6474
10.5665 | -104.6495
10.5703 | -104.6556
10.574 | -104.6665
10.5777 | -104.6798
10.5814 | -104.6948
10.5852 | -104.7036
10.5889 | -104.7157
10.5926 | -104.7307
10.5963 | -104.7435
10.6001 | -104.7591
10.6038 | -104.7768
10.6075 | -104.7975
10.6113 | -104.8154
10.615 | -104.838
10.6187 | -104.8625
10.6224 | -104.8901
10.6262 | -104.9151
10.6299 | -104.9378
10.6336 | -104.9635
10.6373 | -104.9902
10.6411 | -105.0157
10.6448 | -105.0383
10.6485 | -105.0594
10.6523 | -105.0764
10.656 | -105.088
10.6597 | -105.0995
10.6634 | -105.1065
10.6672 | -105.1086
10.6709 | -105.1009
10.6746 | -105.0877
10.6783 | -105.0715
10.6821 | -105.0529
10.6858 | -105.0317
10.6895 | -105.0067
10.6933 | -104.9818
10.697 | -104.9522
10.7007 | -104.9211
10.7044 | -104.8933
10.7082 | -104.8638
10.7119 | -104.8393
10.7156 | -104.8203
10.7193 | -104.7978
10.7231 | -104.7771
10.7268 | -104.7594
10.7305 | -104.7426
10.7343 | -104.731
10.738 | -104.7263
10.7417 | -104.7184
10.7454 | -104.7112
10.7492 | -104.7082
10.7529 | -104.7012
10.7566 | -104.6944
10.7603 | -104.69
10.7641 | -104.6859
10.7678 | -104.6821
10.7715 | -104.6743
10.7753 | -104.6671
10.779 | -104.6617
10.7827 | -104.6579
10.7864 | -104.6565
10.7902 | -104.6513
10.7939 | -104.647
10.7976 | -104.6477
10.8013 | -104.6531
10.8051 | -104.6596
10.8088 | -104.6683
10.8125 | -104.6822
10.8163 | -104.6971
10.82 | -104.7116
10.8237 | -104.7294
10.8274 | -104.7525
10.8312 | -104.7749
10.8349 | -104.7997
10.8386 | -104.8218
10.8423 | -104.8431
10.8461 | -104.8687
10.8498 | -104.8918
10.8535 | -104.9105
10.8572 | -104.9263
10.861 | -104.9372
10.8647 | -104.9446
10.8684 | -104.9484
10.8722 | -104.9519
10.8759 | -104.9535
10.8796 | -104.9516
10.8833 | -104.9456
10.8871 | -104.9433
10.8908 | -104.9402
10.8945 | -104.9347
10.8982 | -104.9349
10.902 | -104.9365
10.9057 | -104.933
10.9094 | -104.9337
10.9132 | -104.9359
10.9169 | -104.9372
10.9206 | -104.9336
10.9243 | -104.9327
10.9281 | -104.9354
10.9318 | -104.9392
10.9355 | -104.9425
10.9392 | -104.9437
10.943 | -104.9412
10.9467 | -104.9374
10.9504 | -104.9317
10.9542 | -104.9223
10.9579 | -104.9064
10.9616 | -104.8877
10.9653 | -104.8682
10.9691 | -104.8449
10.9728 | -104.8182
10.9765 | -104.7902
10.9802 | -104.7636
10.984 | -104.7374
10.9877 | -104.7094
10.9914 | -104.6868
10.9952 | -104.6659
10.9989 | -104.6512
11.0026 | -104.6401
11.0063 | -104.6294
11.0101 | -104.6241
11.0138 | -104.628
11.0175 | -104.6343
11.0212 | -104.6383
11.025 | -104.6481
11.0287 | -104.6601
11.0324 | -104.6695
11.0362 | -104.6798
11.0399 | -104.6911
11.0436 | -104.7037
11.0473 | -104.7128
11.0511 | -104.7202
11.0548 | -104.7275
11.0585 | -104.729
11.0622 | -104.7316
11.066 | -104.7392
11.0697 | -104.7476
11.0734 | -104.7565
11.0772 | -104.7667
11.0809 | -104.7823
11.0846 | -104.7996
11.0883 | -104.8153
11.0921 | -104.833
11.0958 | -104.8507
11.0995 | -104.8695
11.1032 | -104.8919
11.107 | -104.9165
11.1107 | -104.9419
11.1144 | -104.9644
11.1182 | -104.9888
11.1219 | -105.0162
11.1256 | -105.0423
11.1293 | -105.0662
11.1331 | -105.0897
11.1368 | -105.1108
11.1405 | -105.1304
11.1442 | -105.1445
11.148 | -105.157
11.1517 | -105.1671
11.1554 | -105.1759
11.1592 | -105.1775
11.1629 | -105.1741
11.1666 | -105.1699
11.1703 | -105.1613
11.1741 | -105.1487
11.1778 | -105.138
11.1815 | -105.1305
11.1852 | -105.122
11.189 | -105.112
11.1927 | -105.1035
11.1964 | -105.0943
11.2001 | -105.0864
11.2039 | -105.0794
11.2076 | -105.073
11.2113 | -105.0695
11.2151 | -105.0685
11.2188 | -105.0717
11.2225 | -105.0738
11.2262 | -105.0735
11.23 | -105.0765
11.2337 | -105.0809
11.2374 | -105.0831
11.2411 | -105.0849
11.2449 | -105.0838
11.2486 | -105.0848
11.2523 | -105.0899
11.2561 | -105.0957
11.2598 | -105.1006
11.2635 | -105.1027
11.2672 | -105.1061
11.271 | -105.1104
11.2747 | -105.1164
11.2784 | -105.1243
11.2821 | -105.1302
11.2859 | -105.1389
11.2896 | -105.1598
11.2933 | -105.1735
11.2971 | -105.1841
11.3008 | -105.1983
11.3045 | -105.217
11.3082 | -105.2396
11.312 | -105.2606
11.3157 | -105.28
11.3194 | -105.3017
11.3231 | -105.3254
11.3269 | -105.3506
11.3306 | -105.3727
11.3343 | -105.3922
11.3381 | -105.4098
11.3418 | -105.4195
11.3455 | -105.4273
11.3492 | -105.4325
11.353 | -105.4332
11.3567 | -105.4333
11.3604 | -105.4358
11.3641 | -105.435
11.3679 | -105.4298
11.3716 | -105.4251
11.3753 | -105.4199
11.3791 | -105.4138
11.3828 | -105.4024
11.3865 | -105.3935
11.3902 | -105.3812
11.394 | -105.3678
11.3977 | -105.3528
11.4014 | -105.3452
11.4051 | -105.3362
11.4089 | -105.3215
11.4126 | -105.3082
11.4163 | -105.2918
11.4201 | -105.2755
11.4238 | -105.2586
11.4275 | -105.2447
11.4312 | -105.2284
11.435 | -105.2079
11.4387 | -105.1863
11.4424 | -105.1634
11.4461 | -105.1405
11.4499 | -105.1189
11.4536 | -105.0987
11.4573 | -105.0771
11.4611 | -105.055
11.4648 | -105.0388
11.4685 | -105.0228
11.4722 | -105.0059
11.476 | -104.9923
11.4797 | -104.9839
11.4834 | -104.9787
11.4871 | -104.9747
11.4909 | -104.9715
11.4946 | -104.9707
11.4983 | -104.973
11.5021 | -104.9755
11.5058 | -104.9768
11.5095 | -104.9834
11.5132 | -104.9915
11.517 | -104.9963
11.5207 | -105.004
11.5244 | -105.0118
11.5281 | -105.0135
11.5319 | -105.0128
11.5356 | -105.0161
11.5393 | -105.0175
11.543 | -105.0152
11.5468 | -105.0095
11.5505 | -105.0026
11.5542 | -104.9907
11.558 | -104.9803
11.5617 | -104.9652
11.5654 | -104.9529
11.5691 | -104.938
11.5729 | -104.9264
11.5766 | -104.9193
11.5803 | -104.9085
11.584 | -104.9009
11.5878 | -104.8966
11.5915 | -104.8951
11.5952 | -104.8938
11.599 | -104.8954
11.6027 | -104.9065
11.6064 | -104.9235
11.6101 | -104.9394
11.6139 | -104.959
11.6176 | -104.9809
11.6213 | -105.005
11.625 | -105.0362
11.6288 | -105.0678
11.6325 | -105.0962
11.6362 | -105.1235
11.64 | -105.1505
11.6437 | -105.1747
11.6474 | -105.2001
11.6511 | -105.2238
11.6549 | -105.2479
11.6586 | -105.2736
11.6623 | -105.2981
11.666 | -105.321
11.6698 | -105.3427
11.6735 | -105.362
11.6772 | -105.3823
11.681 | -105.4003
11.6847 | -105.4191
11.6884 | -105.4335
11.6921 | -105.4491
11.6959 | -105.4671
11.6996 | -105.4808
11.7033 | -105.4859
11.707 | -105.4919
11.7108 | -105.4976
11.7145 | -105.5017
11.7182 | -105.5024
11.722 | -105.5019
11.7257 | -105.4963
11.7294 | -105.4869
11.7331 | -105.4701
11.7369 | -105.4488
11.7406 | -105.4249
11.7443 | -105.4027
11.748 | -105.3835
11.7518 | -105.3617
11.7555 | -105.3357
11.7592 | -105.3115
11.763 | -105.29
11.7667 | -105.2711
11.7704 | -105.2515
11.7741 | -105.2343
11.7779 | -105.221
11.7816 | -105.2056
11.7853 | -105.1931
11.789 | -105.1838
11.7928 | -105.1745
11.7965 | -105.1672
11.8002 | -105.1609
11.804 | -105.1581
11.8077 | -105.1536
11.8114 | -105.1486
11.8151 | -105.1388
11.8189 | -105.1297
11.8226 | -105.1196
11.8263 | -105.1049
11.83 | -105.0931
11.8338 | -105.082
11.8375 | -105.0731
11.8412 | -105.0656
11.845 | -105.0592
11.8487 | -105.049
11.8524 | -105.0409
11.8561 | -105.0353
11.8599 | -105.0344
11.8636 | -105.0328
11.8673 | -105.0338
11.871 | -105.0391
11.8748 | -105.0461
11.8785 | -105.051
11.8822 | -105.0603
11.8859 | -105.0715
11.8897 | -105.0822
11.8934 | -105.0941
11.8971 | -105.1068
11.9009 | -105.1212
11.9046 | -105.1347
11.9083 | -105.1513
11.912 | -105.1652
11.9158 | -105.1768
11.9195 | -105.1858
11.9232 | -105.1967
11.9269 | -105.2105
11.9307 | -105.2201
11.9344 | -105.2289
11.9381 | -105.2391
11.9419 | -105.2467
11.9456 | -105.2496
11.9493 | -105.2544
11.953 | -105.2607
11.9568 | -105.2659
11.9605 | -105.269
11.9642 | -105.2713
11.9679 | -105.2739
11.9717 | -105.2785
11.9754 | -105.2829
11.9791 | -105.2886
11.9829 | -105.2945
11.9866 | -105.3003
11.9903 | -105.3049
11.994 | -105.311
11.9978 | -105.3154
12.0015 | -105.3199
12.0052 | -105.3257
12.0089 | -105.3279
12.0127 | -105.3284
12.0164 | -105.3288
12.0201 | -105.3317
12.0239 | -105.3363
12.0276 | -105.3374
12.0313 | -105.3361
12.035 | -105.3323
12.0388 | -105.3266
12.0425 | -105.3152
12.0462 | -105.3041
12.0499 | -105.2928
12.0537 | -105.2789
12.0574 | -105.2604
12.0611 | -105.2426
12.0649 | -105.2271
12.0686 | -105.2113
12.0723 | -105.1978
12.076 | -105.1849
12.0798 | -105.1761
12.0835 | -105.168
12.0872 | -105.1612
12.0909 | -105.1573
12.0947 | -105.1581
12.0984 | -105.1597
12.1021 | -105.1651
12.1059 | -105.1736
12.1096 | -105.1846
12.1133 | -105.1954
12.117 | -105.2076
12.1208 | -105.2265
12.1245 | -105.2454
12.1282 | -105.2654
12.1319 | -105.2857
12.1357 | -105.307
12.1394 | -105.3276
12.1431 | -105.3425
12.1469 | -105.3556
12.1506 | -105.3668
12.1543 | -105.3758
12.158 | -105.3824
12.1618 | -105.3934
12.1655 | -105.4037
12.1692 | -105.4086
12.1729 | -105.4138
12.1767 | -105.4205
12.1804 | -105.4176
12.1841 | -105.4174
12.1878 | -105.4187
12.1916 | -105.4182
12.1953 | -105.4126
12.199 | -105.4064
12.2028 | -105.4031
12.2065 | -105.4001
12.2102 | -105.4003
12.2139 | -105.4014
12.2177 | -105.4005
12.2214 | -105.4021
12.2251 | -105.4037
12.2288 | -105.4047
12.2326 | -105.4027
12.2363 | -105.4004
12.24 | -105.3983
12.2438 | -105.3951
12.2475 | -105.3894
12.2512 | -105.3835
12.2549 | -105.3755
12.2587 | -105.3656
12.2624 | -105.3574
12.2661 | -105.3494
12.2698 | -105.3397
12.2736 | -105.3297
12.2773 | -105.3167
12.281 | -105.3056
12.2848 | -105.2957
12.2885 | -105.285
12.2922 | -105.2754
12.2959 | -105.2652
12.2997 | -105.2567
12.3034 | -105.2473
12.3071 | -105.2397
12.3108 | -105.2299
12.3146 | -105.2236
12.3183 | -105.2247
12.322 | -105.224
12.3258 | -105.2207
12.3295 | -105.2222
12.3332 | -105.2236
12.3369 | -105.2226
12.3407 | -105.2235
12.3444 | -105.2275
12.3481 | -105.2331
12.3518 | -105.2413
12.3556 | -105.2557
12.3593 | -105.2684
12.363 | -105.2772
12.3668 | -105.2921
12.3705 | -105.3054
12.3742 | -105.324
12.3779 | -105.3429
12.3817 | -105.3604
12.3854 | -105.3792
12.3891 | -105.3967
12.3928 | -105.4128
12.3966 | -105.4288
12.4003 | -105.4366
12.404 | -105.447
12.4078 | -105.4592
12.4115 | -105.4721
12.4152 | -105.4812
12.4189 | -105.4887
12.4227 | -105.4964
12.4264 | -105.5034
12.4301 | -105.507
12.4338 | -105.5126
12.4376 | -105.5194
12.4413 | -105.5254
12.445 | -105.5273
12.4488 | -105.5312
12.4525 | -105.536
12.4562 | -105.5404
12.4599 | -105.5441
12.4637 | -105.5455
12.4674 | -105.5433
12.4711 | -105.5418
12.4748 | -105.5419
12.4786 | -105.54
12.4823 | -105.5383
12.486 | -105.5358
12.4898 | -105.5324
12.4935 | -105.5251
12.4972 | -105.5181
12.5009 | -105.5092
12.5047 | -105.5009
12.5084 | -105.4862
12.5121 | -105.4753
12.5158 | -105.4619
12.5196 | -105.4439
12.5233 | -105.4283
12.527 | -105.4131
12.5307 | -105.3976
12.5345 | -105.3802
12.5382 | -105.3625
12.5419 | -105.3474
12.5457 | -105.3331
12.5494 | -105.3203
12.5531 | -105.3125
12.5568 | -105.3016
12.5606 | -105.2895
12.5643 | -105.2805
12.568 | -105.2762
12.5717 | -105.2717
12.5755 | -105.2721
12.5792 | -105.2781
12.5829 | -105.2845
12.5867 | -105.2835
12.5904 | -105.2846
12.5941 | -105.2896
12.5978 | -105.2963
12.6016 | -105.3017
12.6053 | -105.3062
12.609 | -105.3091
12.6127 | -105.312
12.6165 | -105.3098
12.6202 | -105.3051
12.6239 | -105.3016
12.6277 | -105.3019
12.6314 | -105.2989
12.6351 | -105.2944
12.6388 | -105.2893
12.6426 | -105.2869
12.6463 | -105.2835
12.65 | -105.2838
12.6537 | -105.2806
12.6575 | -105.2725
12.6612 | -105.265
12.6649 | -105.2601
12.6687 | -105.2562
12.6724 | -105.2527
12.6761 | -105.2503
12.6798 | -105.248
12.6836 | -105.2478
12.6873 | -105.2477
12.691 | -105.2532
12.6947 | -105.255
12.6985 | -105.2559
12.7022 | -105.2629
12.7059 | -105.268
12.7097 | -105.2731
12.7134 | -105.2807
12.7171 | -105.2908
12.7208 | -105.3022
12.7246 | -105.3136
12.7283 | -105.3224
12.732 | -105.3268
12.7357 | -105.3298
12.7395 | -105.3367
12.7432 | -105.3473
12.7469 | -105.3564
12.7507 | -105.36
12.7544 | -105.3646
12.7581 | -105.3682
12.7618 | -105.3683
12.7656 | -105.3694
12.7693 | -105.3702
12.773 | -105.3659
12.7767 | -105.3575
12.7805 | -105.3495
12.7842 | -105.3399
12.7879 | -105.328
12.7917 | -105.3197
12.7954 | -105.3129
12.7991 | -105.3037
12.8028 | -105.2919
12.8066 | -105.2764
12.8103 | -105.2639
12.814 | -105.2526
12.8177 | -105.2434
12.8215 | -105.2346
12.8252 | -105.2256
12.8289 | -105.2194
12.8326 | -105.2148
12.8364 | -105.2108
12.8401 | -105.2084
12.8438 | -105.2066
12.8476 | -105.2032
12.8513 | -105.2014
12.855 | -105.2033
12.8587 | -105.2052
12.8625 | -105.2036
12.8662 | -105.2067
12.8699 | -105.2118
12.8736 | -105.2148
12.8774 | -105.2196
12.8811 | -105.225
12.8848 | -105.2276
12.8886 | -105.2319
12.8923 | -105.2376
12.896 | -105.2399
12.8997 | -105.2411
12.9035 | -105.2401
12.9072 | -105.2389
12.9109 | -105.2376
12.9146 | -105.2392
12.9184 | -105.242
12.9221 | -105.2429
12.9258 | -105.2425
12.9296 | -105.2414
12.9333 | -105.2414
12.937 | -105.2416
12.9407 | -105.239
12.9445 | -105.2339
12.9482 | -105.2317
12.9519 | -105.2322
12.9556 | -105.2363
12.9594 | -105.2408
12.9631 | -105.2444
12.9668 | -105.2456
12.9706 | -105.2514
12.9743 | -105.2591
12.978 | -105.2633
12.9817 | -105.2722
12.9855 | -105.2864
12.9892 | -105.3038
12.9929 | -105.3225
12.9966 | -105.3365
13.0004 | -105.3545
13.0041 | -105.3728
13.0078 | -105.3933
13.0116 | -105.4123
13.0153 | -105.4286
13.019 | -105.4449
13.0227 | -105.4638
13.0265 | -105.4788
13.0302 | -105.4901
13.0339 | -105.5013
13.0376 | -105.509
13.0414 | -105.5143
13.0451 | -105.5195
13.0488 | -105.5214
13.0526 | -105.5188
13.0563 | -105.5161
13.06 | -105.5133
13.0637 | -105.5098
13.0675 | -105.5045
13.0712 | -105.4983
13.0749 | -105.491
13.0786 | -105.4821
13.0824 | -105.4747
13.0861 | -105.4686
13.0898 | -105.4602
13.0936 | -105.4522
13.0973 | -105.4472
13.101 | -105.4424
13.1047 | -105.4335
13.1085 | -105.4247
13.1122 | -105.4203
13.1159 | -105.4155
13.1196 | -105.4097
13.1234 | -105.406
13.1271 | -105.404
13.1308 | -105.4003
13.1346 | -105.3958
13.1383 | -105.3924
13.142 | -105.3874
13.1457 | -105.3818
13.1495 | -105.3788
13.1532 | -105.3803
13.1569 | -105.3834
13.1606 | -105.3795
13.1644 | -105.3732
13.1681 | -105.3687
13.1718 | -105.3582
13.1755 | -105.3482
13.1793 | -105.3382
13.183 | -105.3294
13.1867 | -105.3206
13.1905 | -105.3111
13.1942 | -105.2995
13.1979 | -105.2877
13.2016 | -105.2772
13.2054 | -105.2693
13.2091 | -105.2609
13.2128 | -105.2516
13.2165 | -105.2446
13.2203 | -105.2401
13.224 | -105.2343
13.2277 | -105.2284
13.2315 | -105.225
13.2352 | -105.2239
13.2389 | -105.2256
13.2426 | -105.2305
13.2464 | -105.2349
13.2501 | -105.2425
13.2538 | -105.248
13.2575 | -105.2596
13.2613 | -105.2727
13.265 | -105.2868
13.2687 | -105.3035
13.2725 | -105.318
13.2762 | -105.3294
13.2799 | -105.3427
13.2836 | -105.3572
13.2874 | -105.3699
13.2911 | -105.3807
13.2948 | -105.3943
13.2985 | -105.4076
13.3023 | -105.4179
13.306 | -105.4288
13.3097 | -105.4364
13.3135 | -105.4411
13.3172 | -105.4489
13.3209 | -105.4582
13.3246 | -105.4675
13.3284 | -105.4709
13.3321 | -105.4734
13.3358 | -105.4769
13.3395 | -105.4762
13.3433 | -105.4735
13.347 | -105.4725
13.3507 | -105.4732
13.3545 | -105.4747
13.3582 | -105.4752
13.3619 | -105.4744
13.3656 | -105.4712
13.3694 | -105.4648
13.3731 | -105.4577
13.3768 | -105.4497
13.3805 | -105.4397
13.3843 | -105.4356
13.388 | -105.4298
13.3917 | -105.4194
13.3955 | -105.4077
13.3992 | -105.3964
13.4029 | -105.3876
13.4066 | -105.3792
13.4104 | -105.3715
13.4141 | -105.3631
13.4178 | -105.356
13.4215 | -105.3476
13.4253 | -105.339
13.429 | -105.3329
13.4327 | -105.3231
13.4365 | -105.3161
13.4402 | -105.3165
13.4439 | -105.3136
13.4476 | -105.3119
13.4514 | -105.3098
13.4551 | -105.3109
13.4588 | -105.3109
13.4625 | -105.3108
13.4663 | -105.3159
13.47 | -105.3233
13.4737 | -105.3313
13.4775 | -105.3381
13.4812 | -105.3452
13.4849 | -105.3516
13.4886 | -105.3579
13.4924 | -105.3626
13.4961 | -105.3653
13.4998 | -105.3686
13.5035 | -105.3744
13.5073 | -105.3797
13.511 | -105.3826
13.5147 | -105.3869
13.5184 | -105.3909
13.5222 | -105.3941
13.5259 | -105.3965
13.5296 | -105.3963
13.5334 | -105.3983
13.5371 | -105.3984
13.5408 | -105.397
13.5445 | -105.3947
13.5483 | -105.3936
13.552 | -105.3943
13.5557 | -105.3914
13.5594 | -105.3879
13.5632 | -105.3875
13.5669 | -105.3861
13.5706 | -105.3871
13.5744 | -105.3876
13.5781 | -105.3864
13.5818 | -105.3885
13.5855 | -105.3921
13.5893 | -105.3949
13.593 | -105.3983
13.5967 | -105.3996
13.6004 | -105.4021
13.6042 | -105.4056
13.6079 | -105.4123
13.6116 | -105.4192
13.6154 | -105.4263
13.6191 | -105.4353
13.6228 | -105.4453
13.6265 | -105.4616
13.6303 | -105.4791
13.634 | -105.4946
13.6377 | -105.5109
13.6414 | -105.5277
13.6452 | -105.5507
13.6489 | -105.5728
13.6526 | -105.5942
13.6564 | -105.6189
13.6601 | -105.6461
13.6638 | -105.6723
13.6675 | -105.6999
13.6713 | -105.7242
13.675 | -105.7477
13.6787 | -105.7745
13.6824 | -105.8026
13.6862 | -105.8254
13.6899 | -105.8452
13.6936 | -105.8689
13.6974 | -105.8909
13.7011 | -105.9084
13.7048 | -105.9211
13.7085 | -105.9315
13.7123 | -105.9383
13.716 | -105.9443
13.7197 | -105.9488
13.7234 | -105.9477
13.7272 | -105.944
13.7309 | -105.941
13.7346 | -105.9364
13.7384 | -105.9303
13.7421 | -105.9238
13.7458 | -105.9128
13.7495 | -105.8979
13.7533 | -105.8879
13.757 | -105.8742
13.7607 | -105.8569
13.7644 | -105.8426
13.7682 | -105.8268
13.7719 | -105.811
13.7756 | -105.7957
13.7794 | -105.7814
13.7831 | -105.7713
13.7868 | -105.763
13.7905 | -105.755
13.7943 | -105.7505
13.798 | -105.746
13.8017 | -105.7437
13.8054 | -105.7435
13.8092 | -105.7459
13.8129 | -105.7516
13.8166 | -105.7554
13.8203 | -105.7616
13.8241 | -105.7649
13.8278 | -105.7715
13.8315 | -105.7804
13.8353 | -105.7911
13.839 | -105.8012
13.8427 | -105.8072
13.8464 | -105.8126
13.8502 | -105.8194
13.8539 | -105.8235
13.8576 | -105.8262
13.8613 | -105.8282
13.8651 | -105.8282
13.8688 | -105.8313
13.8725 | -105.8306
13.8763 | -105.8252
13.88 | -105.8183
13.8837 | -105.8147
13.8874 | -105.8105
13.8912 | -105.801
13.8949 | -105.7916
13.8986 | -105.7842
13.9023 | -105.779
13.9061 | -105.7727
13.9098 | -105.7664
13.9135 | -105.7596
13.9173 | -105.7568
13.921 | -105.7573
13.9247 | -105.7597
13.9284 | -105.7574
13.9322 | -105.7555
13.9359 | -105.7572
13.9396 | -105.7597
13.9433 | -105.761
13.9471 | -105.7653
13.9508 | -105.7693
13.9545 | -105.7754
13.9583 | -105.7827
13.962 | -105.7819
13.9657 | -105.7782
13.9694 | -105.7791
13.9732 | -105.7791
13.9769 | -105.7705
13.9806 | -105.7665
13.9843 | -105.7613
13.9881 | -105.7554
13.9918 | -105.7501
13.9955 | -105.7416
13.9993 | -105.7338
14.003 | -105.7237
14.0067 | -105.7105
14.0104 | -105.6992
14.0142 | -105.6858
14.0179 | -105.6706
14.0216 | -105.6544
14.0253 | -105.6367
14.0291 | -105.6171
14.0328 | -105.6015
14.0365 | -105.5861
14.0403 | -105.5694
14.044 | -105.5561
14.0477 | -105.5459
14.0514 | -105.5342
14.0552 | -105.5269
14.0589 | -105.5216
14.0626 | -105.5156
14.0663 | -105.5125
14.0701 | -105.5109
14.0738 | -105.516
14.0775 | -105.5231
14.0813 | -105.53
14.085 | -105.5403
14.0887 | -105.5531
14.0924 | -105.5675
14.0962 | -105.5788
14.0999 | -105.59
14.1036 | -105.6021
14.1073 | -105.6151
14.1111 | -105.6294
14.1148 | -105.6395
14.1185 | -105.6495
14.1223 | -105.6575
14.126 | -105.6606
14.1297 | -105.6663
14.1334 | -105.6706
14.1372 | -105.6697
14.1409 | -105.6657
14.1446 | -105.6625
14.1483 | -105.6585
14.1521 | -105.6514
14.1558 | -105.6461
14.1595 | -105.6412
14.1632 | -105.6404
14.167 | -105.6392
14.1707 | -105.6321
14.1744 | -105.6233
14.1782 | -105.615
14.1819 | -105.6071
14.1856 | -105.599
14.1893 | -105.5905
14.1931 | -105.5816
14.1968 | -105.5704
14.2005 | -105.5597
14.2042 | -105.5495
14.208 | -105.5411
14.2117 | -105.5358
14.2154 | -105.527
14.2192 | -105.5196
14.2229 | -105.5155
14.2266 | -105.5089
14.2303 | -105.5043
14.2341 | -105.5003
14.2378 | -105.4937
14.2415 | -105.4868
14.2452 | -105.4829
14.249 | -105.4782
14.2527 | -105.4745
14.2564 | -105.4699
14.2602 | -105.4664
14.2639 | -105.4666
14.2676 | -105.47
14.2713 | -105.4727
14.2751 | -105.4737
14.2788 | -105.4744
14.2825 | -105.4776
14.2862 | -105.4818
14.29 | -105.4837
14.2937 | -105.489
14.2974 | -105.4965
14.3012 | -105.5057
14.3049 | -105.5179
14.3086 | -105.5317
14.3123 | -105.547
14.3161 | -105.5632
14.3198 | -105.5819
14.3235 | -105.5981
14.3272 | -105.6184
14.331 | -105.6413
14.3347 | -105.6648
14.3384 | -105.6919
14.3422 | -105.7188
14.3459 | -105.7468
14.3496 | -105.7715
14.3533 | -105.7948
14.3571 | -105.8215
14.3608 | -105.8471
14.3645 | -105.8734
14.3682 | -105.8964
14.372 | -105.9208
14.3757 | -105.9458
14.3794 | -105.9694
14.3832 | -105.9903
14.3869 | -106.0074
14.3906 | -106.0207
14.3943 | -106.0336
14.3981 | -106.0462
14.4018 | -106.0582
14.4055 | -106.0673
14.4092 | -106.0716
14.413 | -106.0704
14.4167 | -106.0672
14.4204 | -106.0623
14.4242 | -106.0535
14.4279 | -106.0428
14.4316 | -106.0344
14.4353 | -106.0226
14.4391 | -106.006
14.4428 | -105.9902
14.4465 | -105.9739
14.4502 | -105.9517
14.454 | -105.9268
14.4577 | -105.8997
14.4614 | -105.8747
14.4652 | -105.8493
14.4689 | -105.8164
14.4726 | -105.7811
14.4763 | -105.7488
14.4801 | -105.7179
14.4838 | -105.6803
14.4875 | -105.6423
14.4912 | -105.6073
14.495 | -105.5751
14.4987 | -105.5419
14.5024 | -105.5094
14.5061 | -105.4799
14.5099 | -105.4493
14.5136 | -105.4189
14.5173 | -105.3989
14.5211 | -105.372
14.5248 | -105.3502
14.5285 | -105.3328
14.5322 | -105.3168
14.536 | -105.3004
14.5397 | -105.2874
14.5434 | -105.2788
14.5471 | -105.2712
14.5509 | -105.2676
14.5546 | -105.2673
14.5583 | -105.2709
14.5621 | -105.2766
14.5658 | -105.2839
14.5695 | -105.2941
14.5732 | -105.3063
14.577 | -105.3239
14.5807 | -105.3435
14.5844 | -105.3653
14.5881 | -105.3884
14.5919 | -105.4155
14.5956 | -105.4463
14.5993 | -105.4806
14.6031 | -105.515
14.6068 | -105.5482
14.6105 | -105.584
14.6142 | -105.6211
14.618 | -105.6588
14.6217 | -105.6965
14.6254 | -105.7341
14.6291 | -105.7716
14.6329 | -105.8077
14.6366 | -105.8385
14.6403 | -105.865
14.6441 | -105.8946
14.6478 | -105.9231
14.6515 | -105.9466
14.6552 | -105.9678
14.659 | -105.9888
14.6627 | -106.0076
14.6664 | -106.0212
14.6701 | -106.0287
14.6739 | -106.0303
14.6776 | -106.0328
14.6813 | -106.032
14.6851 | -106.0287
14.6888 | -106.0211
14.6925 | -106.0159
14.6962 | -106.0055
14.7 | -105.9902
14.7037 | -105.9709
14.7074 | -105.9498
14.7111 | -105.9278
14.7149 | -105.9049
14.7186 | -105.8831
14.7223 | -105.862
14.7261 | -105.8411
14.7298 | -105.8222
14.7335 | -105.8027
14.7372 | -105.7802
14.741 | -105.759
14.7447 | -105.7417
14.7484 | -105.7253
14.7521 | -105.7087
14.7559 | -105.69
14.7596 | -105.6732
14.7633 | -105.6603
14.7671 | -105.6511
14.7708 | -105.6425
14.7745 | -105.6304
14.7782 | -105.6175
14.782 | -105.6044
14.7857 | -105.5937
14.7894 | -105.587
14.7931 | -105.58
14.7969 | -105.575
14.8006 | -105.5724
14.8043 | -105.5698
14.8081 | -105.5701
14.8118 | -105.566
14.8155 | -105.5624
14.8192 | -105.5615
14.823 | -105.5577
14.8267 | -105.5566
14.8304 | -105.5567
14.8341 | -105.5585
14.8379 | -105.561
14.8416 | -105.5638
14.8453 | -105.5656
14.8491 | -105.5706
14.8528 | -105.5723
14.8565 | -105.5692
14.8602 | -105.5668
14.864 | -105.5669
14.8677 | -105.5668
14.8714 | -105.5648
14.8751 | -105.5628
14.8789 | -105.5611
14.8826 | -105.5601
14.8863 | -105.5605
14.89 | -105.5593
14.8938 | -105.5567
14.8975 | -105.5502
14.9012 | -105.547
14.905 | -105.5453
14.9087 | -105.5436
14.9124 | -105.5421
14.9161 | -105.5404
14.9199 | -105.5368
14.9236 | -105.5352
14.9273 | -105.5315
14.931 | -105.5274
14.9348 | -105.5255
14.9385 | -105.518
14.9422 | -105.5103
14.946 | -105.5041
14.9497 | -105.4968
14.9534 | -105.4884
14.9571 | -105.48
14.9609 | -105.4722
14.9646 | -105.4641
14.9683 | -105.4559
14.972 | -105.4461
14.9758 | -105.4342
14.9795 | -105.422
14.9832 | -105.4094
14.987 | -105.4029
14.9907 | -105.3974
14.9944 | -105.3913
14.9981 | -105.3847
15.0019 | -105.3786
15.0056 | -105.375
15.0093 | -105.3713
15.013 | -105.3689
15.0168 | -105.3675
15.0205 | -105.3673
15.0242 | -105.3716
15.028 | -105.3746
15.0317 | -105.377
15.0354 | -105.3827
15.0391 | -105.3861
15.0429 | -105.3862
15.0466 | -105.3862
15.0503 | -105.3836
15.054 | -105.3802
15.0578 | -105.3791
15.0615 | -105.3807
15.0652 | -105.3811
15.069 | -105.3835
15.0727 | -105.3847
15.0764 | -105.3847
15.0801 | -105.384
15.0839 | -105.3829
15.0876 | -105.3833
15.0913 | -105.385
15.095 | -105.3861
15.0988 | -105.3875
15.1025 | -105.3895
15.1062 | -105.3939
15.11 | -105.4015
15.1137 | -105.4081
15.1174 | -105.4131
15.1211 | -105.419
15.1249 | -105.4274
15.1286 | -105.4343
15.1323 | -105.4409
15.136 | -105.4537
15.1398 | -105.4666
15.1435 | -105.4807
15.1472 | -105.4944
15.1509 | -105.5068
15.1547 | -105.5264
15.1584 | -105.5425
15.1621 | -105.5567
15.1659 | -105.5724
15.1696 | -105.5889
15.1733 | -105.6022
15.177 | -105.6163
15.1808 | -105.6304
15.1845 | -105.6432
15.1882 | -105.6599
15.1919 | -105.6763
15.1957 | -105.6908
15.1994 | -105.7082
15.2031 | -105.7226
15.2069 | -105.7373
15.2106 | -105.751
15.2143 | -105.7643
15.218 | -105.7751
15.2218 | -105.7849
15.2255 | -105.7961
15.2292 | -105.8068
15.2329 | -105.815
15.2367 | -105.8212
15.2404 | -105.83
15.2441 | -105.8394
15.2479 | -105.8444
15.2516 | -105.8468
15.2553 | -105.8519
15.259 | -105.8588
15.2628 | -105.865
15.2665 | -105.87
15.2702 | -105.8727
15.2739 | -105.8725
15.2777 | -105.8703
15.2814 | -105.8689
15.2851 | -105.867
15.2889 | -105.8654
15.2926 | -105.8639
15.2963 | -105.8611
15.3 | -105.8582
15.3038 | -105.8548
15.3075 | -105.8526
15.3112 | -105.8498
15.3149 | -105.8465
15.3187 | -105.8439
15.3224 | -105.842
15.3261 | -105.8401
15.3299 | -105.8377
15.3336 | -105.8372
15.3373 | -105.8345
15.341 | -105.8313
15.3448 | -105.8299
15.3485 | -105.8301
15.3522 | -105.8312
15.3559 | -105.8286
15.3597 | -105.8267
15.3634 | -105.8256
15.3671 | -105.8271
15.3709 | -105.8279
15.3746 | -105.8275
15.3783 | -105.8262
15.382 | -105.829
15.3858 | -105.8286
15.3895 | -105.8256
15.3932 | -105.8263
15.3969 | -105.8264
15.4007 | -105.8269
15.4044 | -105.828
15.4081 | -105.8253
15.4119 | -105.8232
15.4156 | -105.8235
15.4193 | -105.8218
15.423 | -105.8182
15.4268 | -105.8141
15.4305 | -105.8124
15.4342 | -105.8108
15.4379 | -105.8076
15.4417 | -105.8042
15.4454 | -105.8001
15.4491 | -105.7938
15.4529 | -105.7884
15.4566 | -105.7826
15.4603 | -105.7743
15.464 | -105.7676
15.4678 | -105.7655
15.4715 | -105.7596
15.4752 | -105.7526
15.4789 | -105.7447
15.4827 | -105.7359
15.4864 | -105.7253
15.4901 | -105.7151
15.4939 | -105.7075
15.4976 | -105.6994
15.5013 | -105.6884
15.505 | -105.6773
15.5088 | -105.6668
15.5125 | -105.6588
15.5162 | -105.6499
15.5199 | -105.64
15.5237 | -105.6314
15.5274 | -105.6193
15.5311 | -105.6073
15.5348 | -105.5977
15.5386 | -105.5877
15.5423 | -105.5779
15.546 | -105.5693
15.5498 | -105.5613
15.5535 | -105.5542
15.5572 | -105.5501
15.5609 | -105.5448
15.5647 | -105.5376
15.5684 | -105.5324
15.5721 | -105.5293
15.5758 | -105.5264
15.5796 | -105.5211
15.5833 | -105.5174
15.587 | -105.5147
15.5908 | -105.5141
15.5945 | -105.5143
15.5982 | -105.5147
15.6019 | -105.5168
15.6057 | -105.5169
15.6094 | -105.5176
15.6131 | -105.5213
15.6168 | -105.524
15.6206 | -105.5261
15.6243 | -105.5267
15.628 | -105.5273
15.6318 | -105.5288
15.6355 | -105.5322
15.6392 | -105.5357
15.6429 | -105.5392
15.6467 | -105.5422
15.6504 | -105.546
15.6541 | -105.5494
15.6578 | -105.5524
15.6616 | -105.559
15.6653 | -105.5657
15.669 | -105.5696
15.6728 | -105.5753
15.6765 | -105.5831
15.6802 | -105.5865
15.6839 | -105.5875
15.6877 | -105.5919
15.6914 | -105.5928
15.6951 | -105.5945
15.6988 | -105.5951
15.7026 | -105.5984
15.7063 | -105.6009
15.71 | -105.6032
15.7138 | -105.6041
15.7175 | -105.6047
15.7212 | -105.605
15.7249 | -105.6071
15.7287 | -105.6093
15.7324 | -105.6117
15.7361 | -105.6108
15.7398 | -105.6114
15.7436 | -105.6136
15.7473 | -105.6147
15.751 | -105.6186
15.7548 | -105.6228
15.7585 | -105.6259
15.7622 | -105.6288
15.7659 | -105.6306
15.7697 | -105.6335
15.7734 | -105.6354
15.7771 | -105.6368
15.7808 | -105.6404
15.7846 | -105.6465
15.7883 | -105.6523
15.792 | -105.6576
15.7958 | -105.6607
15.7995 | -105.6646
15.8032 | -105.6682
15.8069 | -105.673
15.8107 | -105.6767
15.8144 | -105.6817
15.8181 | -105.6847
15.8218 | -105.6837
15.8256 | -105.6833
15.8293 | -105.6845
15.833 | -105.6856
15.8368 | -105.6865
15.8405 | -105.6834
15.8442 | -105.6795
15.8479 | -105.6753
15.8517 | -105.6745
15.8554 | -105.6726
15.8591 | -105.6674
15.8628 | -105.6617
15.8666 | -105.6544
15.8703 | -105.647
15.874 | -105.6432
15.8777 | -105.6364
15.8815 | -105.6284
15.8852 | -105.622
15.8889 | -105.6146
15.8927 | -105.609
15.8964 | -105.6056
15.9001 | -105.6008
15.9038 | -105.5953
15.9076 | -105.5954
15.9113 | -105.5974
15.915 | -105.5934
15.9187 | -105.5962
15.9225 | -105.6021
15.9262 | -105.606
15.9299 | -105.6106
15.9337 | -105.6186
15.9374 | -105.6282
15.9411 | -105.6395
15.9448 | -105.6493
15.9486 | -105.6616
15.9523 | -105.6747
15.956 | -105.6871
15.9597 | -105.7003
15.9635 | -105.7157
15.9672 | -105.7317
15.9709 | -105.7457
15.9747 | -105.7607
15.9784 | -105.7759
15.9821 | -105.788
15.9858 | -105.8006
15.9896 | -105.8123
15.9933 | -105.8255
15.997 | -105.8385
16.0007 | -105.853
16.0045 | -105.8671
16.0082 | -105.8791
16.0119 | -105.8899
16.0157 | -105.9009
16.0194 | -105.9111
16.0231 | -105.9199
16.0268 | -105.9273
16.0306 | -105.9364
16.0343 | -105.9435
16.038 | -105.9478
16.0417 | -105.9511
16.0455 | -105.9552
16.0492 | -105.9599
16.0529 | -105.9623
16.0567 | -105.9645
16.0604 | -105.9636
16.0641 | -105.9634
16.0678 | -105.9632
16.0716 | -105.9615
16.0753 | -105.9596
16.079 | -105.9561
16.0827 | -105.9504
16.0865 | -105.9458
16.0902 | -105.9383
16.0939 | -105.9294
16.0977 | -105.9208
16.1014 | -105.9117
16.1051 | -105.9032
16.1088 | -105.8923
16.1126 | -105.877
16.1163 | -105.8622
16.12 | -105.8487
16.1237 | -105.8359
16.1275 | -105.8231
16.1312 | -105.8074
16.1349 | -105.795
16.1387 | -105.7822
16.1424 | -105.7696
16.1461 | -105.7577
16.1498 | -105.7422
16.1536 | -105.7255
16.1573 | -105.7125
16.161 | -105.7003
16.1647 | -105.6872
16.1685 | -105.676
16.1722 | -105.6641
16.1759 | -105.6537
16.1796 | -105.643
16.1834 | -105.6319
16.1871 | -105.623
16.1908 | -105.616
16.1946 | -105.6088
16.1983 | -105.6021
16.202 | -105.5965
16.2057 | -105.5905
16.2095 | -105.5869
16.2132 | -105.5853
16.2169 | -105.5829
16.2206 | -105.5818
16.2244 | -105.5799
16.2281 | -105.5787
16.2318 | -105.579
16.2356 | -105.5803
16.2393 | -105.5812
16.243 | -105.585
16.2467 | -105.5895
16.2505 | -105.5936
16.2542 | -105.602
16.2579 | -105.6129
16.2616 | -105.621
16.2654 | -105.628
16.2691 | -105.6361
16.2728 | -105.6447
16.2766 | -105.6547
16.2803 | -105.6676
16.284 | -105.6812
16.2877 | -105.6918
16.2915 | -105.7029
16.2952 | -105.7171
16.2989 | -105.7304
16.3026 | -105.7437
16.3064 | -105.7576
16.3101 | -105.7687
16.3138 | -105.7817
16.3176 | -105.7963
16.3213 | -105.8103
16.325 | -105.8238
16.3287 | -105.838
16.3325 | -105.8466
16.3362 | -105.8525
16.3399 | -105.8606
16.3436 | -105.8697
16.3474 | -105.8763
16.3511 | -105.8813
16.3548 | -105.8846
16.3586 | -105.887
16.3623 | -105.8864
16.366 | -105.8844
16.3697 | -105.8808
16.3735 | -105.8753
16.3772 | -105.8695
16.3809 | -105.8659
16.3846 | -105.8615
16.3884 | -105.8585
16.3921 | -105.8529
16.3958 | -105.8454
16.3996 | -105.8398
16.4033 | -105.8326
16.407 | -105.8239
16.4107 | -105.8142
16.4145 | -105.8046
16.4182 | -105.7945
16.4219 | -105.7858
16.4256 | -105.7751
16.4294 | -105.7654
16.4331 | -105.7569
16.4368 | -105.7489
16.4406 | -105.7409
16.4443 | -105.7299
16.448 | -105.7213
16.4517 | -105.7152
16.4555 | -105.7061
16.4592 | -105.6939
16.4629 | -105.6817
16.4666 | -105.6698
16.4704 | -105.6576
16.4741 | -105.6471
16.4778 | -105.6367
16.4816 | -105.6291
16.4853 | -105.6231
16.489 | -105.6137
16.4927 | -105.6052
16.4965 | -105.6015
16.5002 | -105.597
16.5039 | -105.5942
16.5076 | -105.5915
16.5114 | -105.587
16.5151 | -105.5807
16.5188 | -105.5753
16.5225 | -105.5748
16.5263 | -105.575
16.53 | -105.5721
16.5337 | -105.5714
16.5375 | -105.572
16.5412 | -105.5743
16.5449 | -105.579
16.5486 | -105.5838
16.5524 | -105.5873
16.5561 | -105.5928
16.5598 | -105.5997
16.5635 | -105.6044
16.5673 | -105.6123
16.571 | -105.6249
16.5747 | -105.6358
16.5785 | -105.6454
16.5822 | -105.6583
16.5859 | -105.6711
16.5896 | -105.6851
16.5934 | -105.6982
16.5971 | -105.7132
16.6008 | -105.7304
16.6045 | -105.7474
16.6083 | -105.7651
16.612 | -105.7839
16.6157 | -105.8029
16.6195 | -105.8212
16.6232 | -105.8379
16.6269 | -105.8562
16.6306 | -105.8737
16.6344 | -105.8873
16.6381 | -105.9032
16.6418 | -105.919
16.6455 | -105.9323
16.6493 | -105.9414
16.653 | -105.9497
16.6567 | -105.9555
16.6605 | -105.9602
16.6642 | -105.9619
16.6679 | -105.9523
16.6716 | -105.9505
16.6754 | -105.9476
16.6791 | -105.9451
16.6828 | -105.9432
16.6865 | -105.9385
16.6903 | -105.9308
16.694 | -105.924
16.6977 | -105.9206
16.7015 | -105.9168
16.7052 | -105.9108
16.7089 | -105.9043
16.7126 | -105.8967
16.7164 | -105.8874
16.7201 | -105.8764
16.7238 | -105.8663
16.7275 | -105.8579
16.7313 | -105.8522
16.735 | -105.8466
16.7387 | -105.8441
16.7425 | -105.8399
16.7462 | -105.8334
16.7499 | -105.8267
16.7536 | -105.8254
16.7574 | -105.8264
16.7611 | -105.8252
16.7648 | -105.8241
16.7685 | -105.8238
16.7723 | -105.8258
16.776 | -105.8276
16.7797 | -105.8297
16.7835 | -105.8283
16.7872 | -105.8295
16.7909 | -105.8325
16.7946 | -105.8348
16.7984 | -105.8375
16.8021 | -105.8442
16.8058 | -105.8488
16.8095 | -105.8508
16.8133 | -105.8536
16.817 | -105.8576
16.8207 | -105.8604
16.8244 | -105.8622
16.8282 | -105.8656
16.8319 | -105.8737
16.8356 | -105.8794
16.8394 | -105.8826
16.8431 | -105.8866
16.8468 | -105.8912
16.8505 | -105.892
16.8543 | -105.8913
16.858 | -105.8934
16.8617 | -105.8966
16.8654 | -105.9015
16.8692 | -105.9056
16.8729 | -105.9097
16.8766 | -105.9124
16.8804 | -105.9159
16.8841 | -105.9194
16.8878 | -105.9219
16.8915 | -105.9235
16.8953 | -105.9254
16.899 | -105.9297
16.9027 | -105.9335
16.9064 | -105.9349
16.9102 | -105.9339
16.9139 | -105.9325
16.9176 | -105.9304
16.9214 | -105.9283
16.9251 | -105.9269
16.9288 | -105.9249
16.9325 | -105.9234
16.9363 | -105.9212
16.94 | -105.9163
16.9437 | -105.9105
16.9474 | -105.9049
16.9512 | -105.9012
16.9549 | -105.896
16.9586 | -105.8884
16.9624 | -105.8817
16.9661 | -105.8755
16.9698 | -105.8695
16.9735 | -105.8616
16.9773 | -105.8518
16.981 | -105.8437
16.9847 | -105.837
16.9884 | -105.8326
16.9922 | -105.8299
16.9959 | -105.8249
16.9996 | -105.8194
17.0034 | -105.8172
17.0071 | -105.8163
17.0108 | -105.8169
17.0145 | -105.818
17.0183 | -105.821
17.022 | -105.8241
17.0257 | -105.8316
17.0294 | -105.839
17.0332 | -105.8445
17.0369 | -105.8507
17.0406 | -105.8572
17.0444 | -105.8617
17.0481 | -105.8684
17.0518 | -105.8761
17.0555 | -105.8841
17.0593 | -105.8914
17.063 | -105.8982
17.0667 | -105.9053
17.0704 | -105.9114
17.0742 | -105.9142
17.0779 | -105.9177
17.0816 | -105.9227
17.0854 | -105.925
17.0891 | -105.9247
17.0928 | -105.9266
17.0965 | -105.9305
17.1003 | -105.9325
17.104 | -105.9313
17.1077 | -105.9288
17.1114 | -105.9264
17.1152 | -105.9264
17.1189 | -105.9262
17.1226 | -105.9264
17.1264 | -105.9273
17.1301 | -105.9287
17.1338 | -105.931
17.1375 | -105.9299
17.1413 | -105.9274
17.145 | -105.9281
17.1487 | -105.9277
17.1524 | -105.9313
17.1562 | -105.9343
17.1599 | -105.9351
17.1636 | -105.9342
17.1673 | -105.9331
17.1711 | -105.9341
17.1748 | -105.9363
17.1785 | -105.9384
17.1823 | -105.9403
17.186 | -105.9414
17.1897 | -105.9421
17.1934 | -105.9418
17.1972 | -105.9409
17.2009 | -105.9423
17.2046 | -105.9432
17.2083 | -105.9422
17.2121 | -105.9421
17.2158 | -105.9408
17.2195 | -105.9398
17.2233 | -105.9414
17.227 | -105.9384
17.2307 | -105.9376
17.2344 | -105.9391
17.2382 | -105.9382
17.2419 | -105.9369
17.2456 | -105.9385
17.2493 | -105.9385
17.2531 | -105.9396
17.2568 | -105.9401
17.2605 | -105.9391
17.2643 | -105.9389
17.268 | -105.9385
17.2717 | -105.939
17.2754 | -105.9365
17.2792 | -105.9347
17.2829 | -105.9342
17.2866 | -105.9325
17.2903 | -105.9294
17.2941 | -105.9262
17.2978 | -105.9225
17.3015 | -105.918
17.3053 | -105.9161
17.309 | -105.913
17.3127 | -105.9089
17.3164 | -105.9045
17.3202 | -105.9001
17.3239 | -105.8934
17.3276 | -105.8858
17.3313 | -105.8785
17.3351 | -105.8724
17.3388 | -105.8657
17.3425 | -105.8568
17.3463 | -105.8471
17.35 | -105.8425
17.3537 | -105.8382
17.3574 | -105.8329
17.3612 | -105.8262
17.3649 | -105.817
17.3686 | -105.8067
17.3723 | -105.7977
17.3761 | -105.7888
17.3798 | -105.7799
17.3835 | -105.7729
17.3873 | -105.7627
17.391 | -105.7509
17.3947 | -105.7403
17.3984 | -105.7283
17.4022 | -105.718
17.4059 | -105.7091
17.4096 | -105.7008
17.4133 | -105.6919
17.4171 | -105.6815
17.4208 | -105.6682
17.4245 | -105.6562
17.4283 | -105.6451
17.432 | -105.6339
17.4357 | -105.624
17.4394 | -105.6136
17.4432 | -105.6043
17.4469 | -105.5949
17.4506 | -105.5841
17.4543 | -105.5741
17.4581 | -105.567
17.4618 | -105.5569
17.4655 | -105.5487
17.4693 | -105.5402
17.473 | -105.5328
17.4767 | -105.5303
17.4804 | -105.526
17.4842 | -105.5213
17.4879 | -105.519
17.4916 | -105.5171
17.4953 | -105.5133
17.4991 | -105.5116
17.5028 | -105.5124
17.5065 | -105.5155
17.5102 | -105.5186
17.514 | -105.5212
17.5177 | -105.5258
17.5214 | -105.5322
17.5252 | -105.5399
17.5289 | -105.5485
17.5326 | -105.5553
17.5363 | -105.5626
17.5401 | -105.5725
17.5438 | -105.5805
17.5475 | -105.5877
17.5512 | -105.5948
17.555 | -105.5986
17.5587 | -105.6032
17.5624 | -105.6101
17.5662 | -105.6176
17.5699 | -105.6241
17.5736 | -105.6326
17.5773 | -105.641
17.5811 | -105.6489
17.5848 | -105.6567
17.5885 | -105.6627
17.5922 | -105.6708
17.596 | -105.6806
17.5997 | -105.6872
17.6034 | -105.6922
17.6072 | -105.7001
17.6109 | -105.71
17.6146 | -105.7207
17.6183 | -105.7319
17.6221 | -105.7411
17.6258 | -105.7505
17.6295 | -105.7581
17.6332 | -105.7659
17.637 | -105.7756
17.6407 | -105.7867
17.6444 | -105.7962
17.6482 | -105.8044
17.6519 | -105.8103
17.6556 | -105.8177
17.6593 | -105.8244
17.6631 | -105.8286
17.6668 | -105.8342
17.6705 | -105.8432
17.6742 | -105.851
17.678 | -105.8562
17.6817 | -105.8613
17.6854 | -105.8677
17.6892 | -105.8731
17.6929 | -105.877
17.6966 | -105.881
17.7003 | -105.8837
17.7041 | -105.8876
17.7078 | -105.8909
17.7115 | -105.8955
17.7152 | -105.8984
17.719 | -105.9011
17.7227 | -105.9032
17.7264 | -105.9022
17.7302 | -105.9021
17.7339 | -105.904
17.7376 | -105.9039
17.7413 | -105.9031
17.7451 | -105.9025
17.7488 | -105.903
17.7525 | -105.9041
17.7562 | -105.902
17.76 | -105.8979
17.7637 | -105.8959
17.7674 | -105.8945
17.7712 | -105.8931
17.7749 | -105.8893
17.7786 | -105.8868
17.7823 | -105.8831
17.7861 | -105.8805
17.7898 | -105.8783
17.7935 | -105.8752
17.7972 | -105.8741
17.801 | -105.8735
17.8047 | -105.873
17.8084 | -105.8726
17.8122 | -105.8715
17.8159 | -105.8694
17.8196 | -105.8648
17.8233 | -105.8599
17.8271 | -105.8556
17.8308 | -105.8517
17.8345 | -105.8462
17.8382 | -105.8418
17.842 | -105.8397
17.8457 | -105.8346
17.8494 | -105.8285
17.8531 | -105.8218
17.8569 | -105.8156
17.8606 | -105.8116
17.8643 | -105.808
17.8681 | -105.8043
17.8718 | -105.7995
17.8755 | -105.7953
17.8792 | -105.7918
17.883 | -105.7871
17.8867 | -105.7851
17.8904 | -105.7816
17.8941 | -105.777
17.8979 | -105.7759
17.9016 | -105.7734
17.9053 | -105.7703
17.9091 | -105.7669
17.9128 | -105.7643
17.9165 | -105.7643
17.9202 | -105.7643
17.924 | -105.7638
17.9277 | -105.7629
17.9314 | -105.7635
17.9351 | -105.7641
17.9389 | -105.7615
17.9426 | -105.7592
17.9463 | -105.7591
17.9501 | -105.7583
17.9538 | -105.7551
17.9575 | -105.7507
17.9612 | -105.7458
17.965 | -105.7396
17.9687 | -105.7338
17.9724 | -105.728
17.9761 | -105.7216
17.9799 | -105.7161
17.9836 | -105.7091
17.9873 | -105.701
17.9911 | -105.6928
17.9948 | -105.6837
17.9985 | -105.6758
18.0022 | -105.6682
18.006 | -105.6617
18.0097 | -105.6526
18.0134 | -105.6452
18.0171 | -105.6403
18.0209 | -105.6342
18.0246 | -105.6278
18.0283 | -105.6195
18.0321 | -105.6131
18.0358 | -105.6059
18.0395 | -105.5994
18.0432 | -105.5933
18.047 | -105.5874
18.0507 | -105.581
18.0544 | -105.5738
18.0581 | -105.5683
18.0619 | -105.5621
18.0656 | -105.5596
18.0693 | -105.5575
18.0731 | -105.5568
18.0768 | -105.5588
18.0805 | -105.5621
18.0842 | -105.5647
18.088 | -105.5662
18.0917 | -105.5695
18.0954 | -105.5715
18.0991 | -105.5732
18.1029 | -105.5773
18.1066 | -105.5812
18.1103 | -105.5871
18.1141 | -105.5954
18.1178 | -105.5991
18.1215 | -105.6054
18.1252 | -105.6143
18.129 | -105.6232
18.1327 | -105.6313
18.1364 | -105.6402
18.1401 | -105.6504
18.1439 | -105.6572
18.1476 | -105.6648
18.1513 | -105.6726
18.155 | -105.6787
18.1588 | -105.6845
18.1625 | -105.6886
18.1662 | -105.6906
18.17 | -105.6934
18.1737 | -105.6957
18.1774 | -105.6954
18.1811 | -105.6947
18.1849 | -105.6932
18.1886 | -105.6916
18.1923 | -105.6901
18.196 | -105.6895
18.1998 | -105.6894
18.2035 | -105.6869
18.2072 | -105.6828
18.211 | -105.6826
18.2147 | -105.6817
18.2184 | -105.6803
18.2221 | -105.6805
18.2259 | -105.6797
18.2296 | -105.6781
18.2333 | -105.6767
18.237 | -105.675
18.2408 | -105.6739
18.2445 | -105.6722
18.2482 | -105.6719
18.252 | -105.6732
18.2557 | -105.6725
18.2594 | -105.672
18.2631 | -105.6715
18.2669 | -105.6728
18.2706 | -105.6728
18.2743 | -105.6714
18.278 | -105.6696
18.2818 | -105.6736
18.2855 | -105.6756
18.2892 | -105.6806
18.293 | -105.6845
18.2967 | -105.6889
18.3004 | -105.6951
18.3041 | -105.699
18.3079 | -105.7047
18.3116 | -105.7127
18.3153 | -105.7231
18.319 | -105.7333
18.3228 | -105.744
18.3265 | -105.753
18.3302 | -105.7636
18.334 | -105.7734
18.3377 | -105.7813
18.3414 | -105.7904
18.3451 | -105.802
18.3489 | -105.8128
18.3526 | -105.8229
18.3563 | -105.8336
18.36 | -105.8435
18.3638 | -105.8532
18.3675 | -105.8634
18.3712 | -105.8727
18.375 | -105.8791
18.3787 | -105.8866
18.3824 | -105.891
18.3861 | -105.8953
18.3899 | -105.8995
18.3936 | -105.903
18.3973 | -105.9003
18.401 | -105.8973
18.4048 | -105.8933
18.4085 | -105.8869
18.4122 | -105.8827
18.416 | -105.8787
18.4197 | -105.8728
18.4234 | -105.8672
18.4271 | -105.8624
18.4309 | -105.8547
18.4346 | -105.8424
18.4383 | -105.8314
18.442 | -105.819
18.4458 | -105.8073
18.4495 | -105.7943
18.4532 | -105.7784
18.457 | -105.7597
18.4607 | -105.7436
18.4644 | -105.7293
18.4681 | -105.7141
18.4719 | -105.6987
18.4756 | -105.6816
18.4793 | -105.6638
18.483 | -105.6463
18.4868 | -105.6292
18.4905 | -105.6119
18.4942 | -105.5956
18.4979 | -105.5817
18.5017 | -105.5677
18.5054 | -105.5527
18.5091 | -105.5394
18.5129 | -105.5266
18.5166 | -105.5146
18.5203 | -105.5074
18.524 | -105.5003
18.5278 | -105.4936
18.5315 | -105.4906
18.5352 | -105.4892
18.5389 | -105.4902
18.5427 | -105.4892
18.5464 | -105.4899
18.5501 | -105.4917
18.5539 | -105.4937
18.5576 | -105.4984
18.5613 | -105.504
18.565 | -105.5088
18.5688 | -105.5154
18.5725 | -105.5226
18.5762 | -105.5289
18.5799 | -105.5331
18.5837 | -105.537
18.5874 | -105.5442
18.5911 | -105.5526
18.5949 | -105.5617
18.5986 | -105.5702
18.6023 | -105.5803
18.606 | -105.5907
18.6098 | -105.5966
18.6135 | -105.603
18.6172 | -105.6123
18.6209 | -105.6218
18.6247 | -105.6285
18.6284 | -105.6364
18.6321 | -105.6427
18.6359 | -105.6467
18.6396 | -105.6514
18.6433 | -105.6587
18.647 | -105.6648
18.6508 | -105.6675
18.6545 | -105.6703
18.6582 | -105.6741
18.6619 | -105.681
18.6657 | -105.6854
18.6694 | -105.6876
18.6731 | -105.6914
18.6769 | -105.692
18.6806 | -105.6902
18.6843 | -105.6889
18.688 | -105.6887
18.6918 | -105.6851
18.6955 | -105.6824
18.6992 | -105.6823
18.7029 | -105.6817
18.7067 | -105.6797
18.7104 | -105.6774
18.7141 | -105.6746
18.7179 | -105.6729
18.7216 | -105.6711
18.7253 | -105.6659
18.729 | -105.6613
18.7328 | -105.6576
18.7365 | -105.6528
18.7402 | -105.6487
18.7439 | -105.6445
18.7477 | -105.6424
18.7514 | -105.6419
18.7551 | -105.6431
18.7589 | -105.6457
18.7626 | -105.6471
18.7663 | -105.6483
18.77 | -105.6501
18.7738 | -105.6505
18.7775 | -105.6494
18.7812 | -105.6512
18.7849 | -105.652
18.7887 | -105.6535
18.7924 | -105.6552
18.7961 | -105.6559
18.7999 | -105.659
18.8036 | -105.6631
18.8073 | -105.6674
18.811 | -105.6707
18.8148 | -105.6744
18.8185 | -105.6799
18.8222 | -105.6847
18.8259 | -105.6878
18.8297 | -105.6934
18.8334 | -105.7001
18.8371 | -105.7054
18.8408 | -105.7115
18.8446 | -105.7188
18.8483 | -105.7275
18.852 | -105.7356
18.8558 | -105.7411
18.8595 | -105.7471
18.8632 | -105.751
18.8669 | -105.7569
18.8707 | -105.7652
18.8744 | -105.7717
18.8781 | -105.7776
18.8818 | -105.7841
18.8856 | -105.7885
18.8893 | -105.7913
18.893 | -105.798
18.8968 | -105.8041
18.9005 | -105.8102
18.9042 | -105.817
18.9079 | -105.8246
18.9117 | -105.8289
18.9154 | -105.8331
18.9191 | -105.8385
18.9228 | -105.8453
18.9266 | -105.8514
18.9303 | -105.8562
18.934 | -105.862
18.9378 | -105.8701
18.9415 | -105.8748
18.9452 | -105.8777
18.9489 | -105.883
18.9527 | -105.889
18.9564 | -105.8922
18.9601 | -105.8974
18.9638 | -105.9026
18.9676 | -105.9078
18.9713 | -105.9124
18.975 | -105.918
18.9788 | -105.9225
18.9825 | -105.9278
18.9862 | -105.9332
18.9899 | -105.9401
18.9937 | -105.9475
18.9974 | -105.9536
19.0011 | -105.9593
19.0048 | -105.9648
19.0086 | -105.9699
19.0123 | -105.9752
19.016 | -105.9781
19.0198 | -105.9804
19.0235 | -105.9806
19.0272 | -105.9802
19.0309 | -105.9815
19.0347 | -105.9846
19.0384 | -105.9885
19.0421 | -105.9923
19.0458 | -105.996
19.0496 | -105.9966
19.0533 | -105.9959
19.057 | -105.9939
19.0608 | -105.9917
19.0645 | -105.9878
19.0682 | -105.9825
19.0719 | -105.9765
19.0757 | -105.9705
19.0794 | -105.9629
19.0831 | -105.9536
19.0868 | -105.9452
19.0906 | -105.9335
19.0943 | -105.9216
19.098 | -105.9117
19.1018 | -105.9006
19.1055 | -105.8891
19.1092 | -105.8792
19.1129 | -105.869
19.1167 | -105.8586
19.1204 | -105.8477
19.1241 | -105.8384
19.1278 | -105.8296
19.1316 | -105.818
19.1353 | -105.8086
19.139 | -105.8005
19.1428 | -105.7881
19.1465 | -105.7773
19.1502 | -105.7716
19.1539 | -105.7668
19.1577 | -105.7623
19.1614 | -105.7587
19.1651 | -105.7536
19.1688 | -105.7481
19.1726 | -105.742
19.1763 | -105.7373
19.18 | -105.7335
19.1837 | -105.7304
19.1875 | -105.7286
19.1912 | -105.7311
19.1949 | -105.7318
19.1987 | -105.7306
19.2024 | -105.7307
19.2061 | -105.7311
19.2098 | -105.7304
19.2136 | -105.7306
19.2173 | -105.7307
19.221 | -105.7302
19.2247 | -105.73
19.2285 | -105.7321
19.2322 | -105.7327
19.2359 | -105.7334
19.2397 | -105.7316
19.2434 | -105.7287
19.2471 | -105.7268
19.2508 | -105.7271
19.2546 | -105.7254
19.2583 | -105.7232
19.262 | -105.7214
19.2657 | -105.7178
19.2695 | -105.715
19.2732 | -105.7123
19.2769 | -105.708
19.2807 | -105.7041
19.2844 | -105.7011
19.2881 | -105.6959
19.2918 | -105.6895
19.2956 | -105.6829
19.2993 | -105.6752
19.303 | -105.6665
19.3067 | -105.6568
19.3105 | -105.6477
19.3142 | -105.6379
19.3179 | -105.631
19.3217 | -105.6239
19.3254 | -105.6154
19.3291 | -105.6075
19.3328 | -105.6011
19.3366 | -105.5934
19.3403 | -105.5856
19.344 | -105.5815
19.3477 | -105.5768
19.3515 | -105.5703
19.3552 | -105.5654
19.3589 | -105.56
19.3627 | -105.5534
19.3664 | -105.5487
19.3701 | -105.5453
19.3738 | -105.5406
19.3776 | -105.5347
19.3813 | -105.5299
19.385 | -105.525
19.3887 | -105.5182
19.3925 | -105.5142
19.3962 | -105.5115
19.3999 | -105.5063
19.4037 | -105.5022
19.4074 | -105.5019
19.4111 | -105.5013
19.4148 | -105.497
19.4186 | -105.4922
19.4223 | -105.4904
19.426 | -105.4883
19.4297 | -105.4845
19.4335 | -105.4782
19.4372 | -105.4735
19.4409 | -105.4688
19.4447 | -105.4642
19.4484 | -105.4614
19.4521 | -105.4552
19.4558 | -105.4457
19.4596 | -105.4378
19.4633 | -105.432
19.467 | -105.425
19.4707 | -105.4175
19.4745 | -105.4123
19.4782 | -105.4044
19.4819 | -105.3965
19.4856 | -105.3903
19.4894 | -105.3843
19.4931 | -105.3786
19.4968 | -105.3718
19.5006 | -105.3648
19.5043 | -105.3582
19.508 | -105.3522
19.5117 | -105.3455
19.5155 | -105.3419
19.5192 | -105.3392
19.5229 | -105.3351
19.5266 | -105.3311
19.5304 | -105.3301
19.5341 | -105.3299
19.5378 | -105.3286
19.5416 | -105.3281
19.5453 | -105.3276
19.549 | -105.3319
19.5527 | -105.3354
19.5565 | -105.3373
19.5602 | -105.3391
19.5639 | -105.3429
19.5676 | -105.3488
19.5714 | -105.3552
19.5751 | -105.3616
19.5788 | -105.3697
19.5826 | -105.3775
19.5863 | -105.3837
19.59 | -105.3924
19.5937 | -105.3983
19.5975 | -105.4056
19.6012 | -105.4158
19.6049 | -105.4252
19.6086 | -105.4333
19.6124 | -105.4435
19.6161 | -105.452
19.6198 | -105.4573
19.6236 | -105.4656
19.6273 | -105.4736
19.631 | -105.4825
19.6347 | -105.4895
19.6385 | -105.4952
19.6422 | -105.5016
19.6459 | -105.5049
19.6496 | -105.5056
19.6534 | -105.5073
19.6571 | -105.5098
19.6608 | -105.5083
19.6646 | -105.5056
19.6683 | -105.5045
19.672 | -105.5009
19.6757 | -105.4983
19.6795 | -105.4946
19.6832 | -105.4888
19.6869 | -105.4815
19.6906 | -105.4729
19.6944 | -105.4657
19.6981 | -105.4601
19.7018 | -105.4499
19.7056 | -105.4364
19.7093 | -105.4227
19.713 | -105.4126
19.7167 | -105.4018
19.7205 | -105.3921
19.7242 | -105.3782
19.7279 | -105.3629
19.7316 | -105.3516
19.7354 | -105.342
19.7391 | -105.3307
19.7428 | -105.3166
19.7466 | -105.3013
19.7503 | -105.2888
19.754 | -105.2761
19.7577 | -105.2666
19.7615 | -105.2565
19.7652 | -105.2463
19.7689 | -105.2385
19.7726 | -105.2299
19.7764 | -105.2225
19.7801 | -105.2146
19.7838 | -105.2067
19.7876 | -105.2016
19.7913 | -105.1975
19.795 | -105.1932
19.7987 | -105.1885
19.8025 | -105.1864
19.8062 | -105.1838
19.8099 | -105.181
19.8136 | -105.1801
19.8174 | -105.1797
19.8211 | -105.1813
19.8248 | -105.1841
19.8285 | -105.184
19.8323 | -105.1892
19.836 | -105.1972
19.8397 | -105.2041
19.8435 | -105.2118
19.8472 | -105.2198
19.8509 | -105.2286
19.8546 | -105.2391
19.8584 | -105.2489
19.8621 | -105.2601
19.8658 | -105.2736
19.8695 | -105.2857
19.8733 | -105.2998
19.877 | -105.3133
19.8807 | -105.326
19.8845 | -105.3414
19.8882 | -105.356
19.8919 | -105.3699
19.8956 | -105.3862
19.8994 | -105.4038
19.9031 | -105.4211
19.9068 | -105.4361
19.9105 | -105.4524
19.9143 | -105.4666
19.918 | -105.4831
19.9217 | -105.4994
19.9255 | -105.5135
19.9292 | -105.5273
19.9329 | -105.5418
19.9366 | -105.5526
19.9404 | -105.562
19.9441 | -105.5738
19.9478 | -105.5856
19.9515 | -105.5944
19.9553 | -105.6013
19.959 | -105.6046
19.9627 | -105.6089
19.9665 | -105.6118
19.9702 | -105.6144
19.9739 | -105.6177
19.9776 | -105.6188
19.9814 | -105.6182
19.9851 | -105.6176
19.9888 | -105.6186
19.9925 | -105.6193
19.9963 | -105.6154
20 | -105.6105
20.0037 | -105.6043
20.0075 | -105.6035
20.0112 | -105.6023
20.0149 | -105.5976
20.0186 | -105.5901
20.0224 | -105.5816
20.0261 | -105.574
20.0298 | -105.5677
20.0335 | -105.5599
20.0373 | -105.5495
20.041 | -105.537
20.0447 | -105.5283
20.0485 | -105.5193
20.0522 | -105.5104
20.0559 | -105.5015
20.0596 | -105.4927
20.0634 | -105.4857
20.0671 | -105.4797
20.0708 | -105.4719
20.0745 | -105.465
20.0783 | -105.459
20.082 | -105.4529
20.0857 | -105.446
20.0895 | -105.4414
20.0932 | -105.4366
20.0969 | -105.4315
20.1006 | -105.428
20.1044 | -105.424
20.1081 | -105.4189
20.1118 | -105.4151
20.1155 | -105.4107
20.1193 | -105.4061
20.123 | -105.402
20.1267 | -105.3996
20.1305 | -105.3967
20.1342 | -105.3941
20.1379 | -105.3907
20.1416 | -105.3881
20.1454 | -105.3857
20.1491 | -105.3835
20.1528 | -105.383
20.1565 | -105.3821
20.1603 | -105.382
20.164 | -105.3806
20.1677 | -105.3778
20.1715 | -105.3747
20.1752 | -105.3703
20.1789 | -105.3676
20.1826 | -105.3633
20.1864 | -105.3571
20.1901 | -105.3515
20.1938 | -105.3473
20.1975 | -105.344
20.2013 | -105.3404
20.205 | -105.3335
20.2087 | -105.3283
20.2124 | -105.3238
20.2162 | -105.3194
20.2199 | -105.3128
20.2236 | -105.3059
20.2274 | -105.2977
20.2311 | -105.2893
20.2348 | -105.2803
20.2385 | -105.2725
20.2423 | -105.264
20.246 | -105.2558
20.2497 | -105.247
20.2534 | -105.2385
20.2572 | -105.2284
20.2609 | -105.2175
20.2646 | -105.2069
20.2684 | -105.1978
20.2721 | -105.1871
20.2758 | -105.1748
20.2795 | -105.1633
20.2833 | -105.153
20.287 | -105.143
20.2907 | -105.1341
20.2944 | -105.1227
20.2982 | -105.1125
20.3019 | -105.1039
20.3056 | -105.0954
20.3094 | -105.0886
20.3131 | -105.0813
20.3168 | -105.0727
20.3205 | -105.0678
20.3243 | -105.0646
20.328 | -105.0619
20.3317 | -105.0565
20.3354 | -105.051
20.3392 | -105.0465
20.3429 | -105.0435
20.3466 | -105.0421
20.3504 | -105.0407
20.3541 | -105.0398
20.3578 | -105.04
20.3615 | -105.0415
20.3653 | -105.0436
20.369 | -105.0482
20.3727 | -105.0516
20.3764 | -105.0547
20.3802 | -105.0593
20.3839 | -105.0639
20.3876 | -105.0687
20.3914 | -105.0749
20.3951 | -105.0824
20.3988 | -105.0884
20.4025 | -105.0951
20.4063 | -105.1037
20.41 | -105.1129
20.4137 | -105.1205
20.4174 | -105.128
20.4212 | -105.1349
20.4249 | -105.1413
20.4286 | -105.1468
20.4324 | -105.1498
20.4361 | -105.1566
20.4398 | -105.1618
20.4435 | -105.1654
20.4473 | -105.1709
20.451 | -105.174
20.4547 | -105.1781
20.4584 | -105.182
20.4622 | -105.1879
20.4659 | -105.1928
20.4696 | -105.1951
20.4734 | -105.1958
20.4771 | -105.1967
20.4808 | -105.1964
20.4845 | -105.1963
20.4883 | -105.1934
20.492 | -105.1889
20.4957 | -105.1887
20.4994 | -105.1878
20.5032 | -105.1861
20.5069 | -105.1823
20.5106 | -105.1777
20.5144 | -105.1772
20.5181 | -105.1779
20.5218 | -105.176
20.5255 | -105.1711
20.5293 | -105.1666
20.533 | -105.1635
20.5367 | -105.1624
20.5404 | -105.1613
20.5442 | -105.1585
20.5479 | -105.1558
20.5516 | -105.1538
20.5553 | -105.1528
20.5591 | -105.1524
20.5628 | -105.1541
20.5665 | -105.1553
20.5703 | -105.1571
20.574 | -105.1601
20.5777 | -105.1597
20.5814 | -105.157
20.5852 | -105.1556
20.5889 | -105.1528
20.5926 | -105.1537
20.5963 | -105.155
20.6001 | -105.155
20.6038 | -105.1534
20.6075 | -105.1519
20.6113 | -105.1483
20.615 | -105.1454
20.6187 | -105.1433
20.6224 | -105.141
20.6262 | -105.1375
20.6299 | -105.1342
20.6336 | -105.1321
20.6373 | -105.1279
20.6411 | -105.122
20.6448 | -105.1151
20.6485 | -105.1101
20.6523 | -105.1064
20.656 | -105.1027
20.6597 | -105.0993
20.6634 | -105.0959
20.6672 | -105.0906
20.6709 | -105.085
20.6746 | -105.0813
20.6783 | -105.0782
20.6821 | -105.0743
20.6858 | -105.072
20.6895 | -105.0699
20.6933 | -105.0676
20.697 | -105.0642
20.7007 | -105.0622
20.7044 | -105.0618
20.7082 | -105.0627
20.7119 | -105.065
20.7156 | -105.0679
20.7193 | -105.0695
20.7231 | -105.0712
20.7268 | -105.0726
20.7305 | -105.0731
20.7343 | -105.0747
20.738 | -105.0782
20.7417 | -105.0819
20.7454 | -105.0866
20.7492 | -105.0904
20.7529 | -105.0954
20.7566 | -105.0998
20.7603 | -105.1063
20.7641 | -105.1122
20.7678 | -105.1184
20.7715 | -105.1262
20.7753 | -105.1342
20.779 | -105.1416
20.7827 | -105.1502
20.7864 | -105.1597
20.7902 | -105.1675
20.7939 | -105.1746
20.7976 | -105.1827
20.8013 | -105.1941
20.8051 | -105.2052
20.8088 | -105.2126
20.8125 | -105.2194
20.8163 | -105.2275
20.82 | -105.2382
20.8237 | -105.2467
20.8274 | -105.254
20.8312 | -105.2603
20.8349 | -105.2671
20.8386 | -105.276
20.8423 | -105.2815
20.8461 | -105.2861
20.8498 | -105.2896
20.8535 | -105.2904
20.8572 | -105.2914
20.861 | -105.2943
20.8647 | -105.2962
20.8684 | -105.2974
20.8722 | -105.3005
20.8759 | -105.3037
20.8796 | -105.3061
20.8833 | -105.3084
20.8871 | -105.3033
20.8908 | -105.2958
20.8945 | -105.2897
20.8982 | -105.2856
20.902 | -105.2816
20.9057 | -105.2782
20.9094 | -105.2744
20.9132 | -105.269
20.9169 | -105.2604
20.9206 | -105.2513
20.9243 | -105.2401
20.9281 | -105.229
20.9318 | -105.219
20.9355 | -105.2086
20.9392 | -105.1969
20.943 | -105.1838
20.9467 | -105.169
20.9504 | -105.1556
20.9542 | -105.142
20.9579 | -105.1298
20.9616 | -105.1183
20.9653 | -105.1075
20.9691 | -105.0941
20.9728 | -105.0811
20.9765 | -105.0669
20.9802 | -105.0526
20.984 | -105.0369
20.9877 | -105.0235
20.9914 | -105.0098
20.9952 | -104.9972
20.9989 | -104.9831
21.0026 | -104.969
21.0063 | -104.9547
21.0101 | -104.9403
21.0138 | -104.9246
21.0175 | -104.9121
21.0212 | -104.9017
21.025 | -104.891
21.0287 | -104.8809
21.0324 | -104.8691
21.0362 | -104.8559
21.0399 | -104.842
21.0436 | -104.8306
21.0473 | -104.8207
21.0511 | -104.81
21.0548 | -104.7987
21.0585 | -104.7867
21.0622 | -104.775
21.066 | -104.7631
21.0697 | -104.7499
21.0734 | -104.7373
21.0772 | -104.7232
21.0809 | -104.7097
21.0846 | -104.696
21.0883 | -104.6801
21.0921 | -104.665
21.0958 | -104.6495
21.0995 | -104.6319
21.1032 | -104.6139
21.107 | -104.5955
21.1107 | -104.5795
21.1144 | -104.5629
21.1182 | -104.5435
21.1219 | -104.5243
21.1256 | -104.5055
21.1293 | -104.4862
21.1331 | -104.4659
21.1368 | -104.4456
21.1405 | -104.427
21.1442 | -104.4082
21.148 | -104.3912
21.1517 | -104.374
21.1554 | -104.357
21.1592 | -104.3404
21.1629 | -104.3229
21.1666 | -104.3053
21.1703 | -104.2897
21.1741 | -104.2739
21.1778 | -104.2587
21.1815 | -104.2451
21.1852 | -104.2355
21.189 | -104.2258
21.1927 | -104.2148
21.1964 | -104.2047
21.2001 | -104.1947
21.2039 | -104.1855
21.2076 | -104.1764
21.2113 | -104.1692
21.2151 | -104.1651
21.2188 | -104.1618
21.2225 | -104.1579
21.2262 | -104.1541
21.23 | -104.1535
21.2337 | -104.1548
21.2374 | -104.1563
21.2411 | -104.1559
21.2449 | -104.156
21.2486 | -104.1567
21.2523 | -104.1555
21.2561 | -104.1549
21.2598 | -104.1561
21.2635 | -104.1571
21.2672 | -104.1596
21.271 | -104.1609
21.2747 | -104.1621
21.2784 | -104.1668
21.2821 | -104.1703
21.2859 | -104.1732
21.2896 | -104.1749
21.2933 | -104.1786
21.2971 | -104.1818
21.3008 | -104.1861
21.3045 | -104.1904
21.3082 | -104.1941
21.312 | -104.1984
21.3157 | -104.2056
21.3194 | -104.2091
21.3231 | -104.2109
21.3269 | -104.2126
21.3306 | -104.2148
21.3343 | -104.2193
21.3381 | -104.2251
21.3418 | -104.229
21.3455 | -104.2311
21.3492 | -104.2331
21.353 | -104.2361
21.3567 | -104.2397
21.3604 | -104.2432
21.3641 | -104.246
21.3679 | -104.2495
21.3716 | -104.2541
21.3753 | -104.2588
21.3791 | -104.2634
21.3828 | -104.2677
21.3865 | -104.2714
21.3902 | -104.2773
21.394 | -104.2827
21.3977 | -104.2871
21.4014 | -104.2923
21.4051 | -104.2996
21.4089 | -104.308
21.4126 | -104.3121
21.4163 | -104.3149
21.4201 | -104.3169
21.4238 | -104.3177
21.4275 | -104.3186
21.4312 | -104.319
21.435 | -104.319
21.4387 | -104.3169
21.4424 | -104.3147
21.4461 | -104.3132
21.4499 | -104.3082
21.4536 | -104.3007
21.4573 | -104.2918
21.4611 | -104.2825
21.4648 | -104.2728
21.4685 | -104.2637
21.4722 | -104.2535
21.476 | -104.2426
21.4797 | -104.2313
21.4834 | -104.2182
21.4871 | -104.202
21.4909 | -104.1857
21.4946 | -104.1699
21.4983 | -104.1532
21.5021 | -104.1353
21.5058 | -104.1135
21.5095 | -104.0951
21.5132 | -104.0764
21.517 | -104.0568
21.5207 | -104.0377
21.5244 | -104.0205
21.5281 | -104.0042
21.5319 | -103.9888
21.5356 | -103.9685
21.5393 | -103.9498
21.543 | -103.9347
21.5468 | -103.922
21.5505 | -103.9098
21.5542 | -103.8982
21.558 | -103.8911
21.5617 | -103.8847
21.5654 | -103.879
21.5691 | -103.8764
21.5729 | -103.8749
21.5766 | -103.8777
21.5803 | -103.8818
21.584 | -103.8861
21.5878 | -103.8929
21.5915 | -103.9039
21.5952 | -103.9172
21.599 | -103.9298
21.6027 | -103.9448
21.6064 | -103.9623
21.6101 | -103.9812
21.6139 | -104.0028
21.6176 | -104.0262
21.6213 | -104.0498
21.625 | -104.076
21.6288 | -104.1042
21.6325 | -104.1335
21.6362 | -104.1632
21.64 | -104.1936
21.6437 | -104.2234
21.6474 | -104.2555
21.6511 | -104.2886
21.6549 | -104.3205
21.6586 | -104.3518
21.6623 | -104.3834
21.666 | -104.4139
21.6698 | -104.4459
21.6735 | -104.4759
21.6772 | -104.5047
21.681 | -104.5316
21.6847 | -104.5594
21.6884 | -104.5867
21.6921 | -104.612
21.6959 | -104.6322
21.6996 | -104.6537
21.7033 | -104.6769
21.707 | -104.6977
21.7108 | -104.7145
21.7145 | -104.7291
21.7182 | -104.7455
21.722 | -104.7616
21.7257 | -104.779
21.7294 | -104.7958
21.7331 | -104.81
21.7369 | -104.821
21.7406 | -104.8276
21.7443 | -104.8358
21.748 | -104.8405
21.7518 | -104.8425
21.7555 | -104.8447
21.7592 | -104.8486
21.763 | -104.8524
21.7667 | -104.8554
21.7704 | -104.8566
21.7741 | -104.8544
21.7779 | -104.8506
21.7816 | -104.8474
21.7853 | -104.8433
21.789 | -104.8381
21.7928 | -104.8349
21.7965 | -104.8318
21.8002 | -104.8275
21.804 | -104.8231
21.8077 | -104.8194
21.8114 | -104.8159
21.8151 | -104.8106
21.8189 | -104.8049
21.8226 | -104.7999
21.8263 | -104.7967
21.83 | -104.7918
21.8338 | -104.7865
21.8375 | -104.7816
21.8412 | -104.7748
21.845 | -104.7734
21.8487 | -104.7712
21.8524 | -104.7679
21.8561 | -104.7652
21.8599 | -104.7633
21.8636 | -104.7611
21.8673 | -104.7586
21.871 | -104.7561
21.8748 | -104.7521
21.8785 | -104.7457
21.8822 | -104.7416
21.8859 | -104.7366
21.8897 | -104.7334
21.8934 | -104.7294
21.8971 | -104.7261
21.9009 | -104.7221
21.9046 | -104.7194
21.9083 | -104.7188
21.912 | -104.7179
21.9158 | -104.7161
21.9195 | -104.7148
21.9232 | -104.7145
21.9269 | -104.7138
21.9307 | -104.715
21.9344 | -104.7148
21.9381 | -104.7152
21.9419 | -104.7159
21.9456 | -104.7171
21.9493 | -104.7183
21.953 | -104.7215
21.9568 | -104.7239
21.9605 | -104.7246
21.9642 | -104.7268
21.9679 | -104.7291
21.9717 | -104.7321
21.9754 | -104.7359
21.9791 | -104.7376
21.9829 | -104.741
21.9866 | -104.7425
21.9903 | -104.7457
21.994 | -104.7496
21.9978 | -104.7531
22.0015 | -104.7582
22.0052 | -104.7642
22.0089 | -104.768
22.0127 | -104.7716
22.0164 | -104.7746
22.0201 | -104.7783
22.0239 | -104.7849
22.0276 | -104.7902
22.0313 | -104.7966
22.035 | -104.8029
22.0388 | -104.8078
22.0425 | -104.8119
22.0462 | -104.8147
22.0499 | -104.8195
22.0537 | -104.8267
22.0574 | -104.8294
22.0611 | -104.8318
22.0649 | -104.8365
22.0686 | -104.8423
22.0723 | -104.8459
22.076 | -104.8477
22.0798 | -104.8481
22.0835 | -104.8491
22.0872 | -104.852
22.0909 | -104.8546
22.0947 | -104.8556
22.0984 | -104.8561
22.1021 | -104.8569
22.1059 | -104.8574
22.1096 | -104.8588
22.1133 | -104.8588
22.117 | -104.8565
22.1208 | -104.8549
22.1245 | -104.8525
22.1282 | -104.8511
22.1319 | -104.8493
22.1357 | -104.8468
22.1394 | -104.8422
22.1431 | -104.8393
22.1469 | -104.8341
22.1506 | -104.8296
22.1543 | -104.8254
22.158 | -104.8176
22.1618 | -104.8106
22.1655 | -104.8046
22.1692 | -104.7972
22.1729 | -104.7898
22.1767 | -104.7826
22.1804 | -104.7755
22.1841 | -104.7691
22.1878 | -104.7638
22.1916 | -104.7593
22.1953 | -104.7537
22.199 | -104.7491
22.2028 | -104.7437
22.2065 | -104.7378
22.2102 | -104.733
22.2139 | -104.7282
22.2177 | -104.7234
22.2214 | -104.7181
22.2251 | -104.7136
22.2288 | -104.7107
22.2326 | -104.7101
22.2363 | -104.7078
22.24 | -104.7064
22.2438 | -104.7068
22.2475 | -104.7102
22.2512 | -104.7111
22.2549 | -104.7132
22.2587 | -104.7171
22.2624 | -104.72
22.2661 | -104.7238
22.2698 | -104.7295
22.2736 | -104.7368
22.2773 | -104.7443
22.281 | -104.7524
22.2848 | -104.7602
22.2885 | -104.7686
22.2922 | -104.7784
22.2959 | -104.7897
22.2997 | -104.8031
22.3034 | -104.8151
22.3071 | -104.8259
22.3108 | -104.8353
22.3146 | -104.8429
22.3183 | -104.8528
22.322 | -104.8636
22.3258 | -104.8737
22.3295 | -104.8858
22.3332 | -104.8992
22.3369 | -104.9091
22.3407 | -104.9192
22.3444 | -104.9292
22.3481 | -104.939
22.3518 | -104.949
22.3556 | -104.9602
22.3593 | -104.9718
22.363 | -104.9829
22.3668 | -104.9914
22.3705 | -104.9992
22.3742 | -105.01
22.3779 | -105.0215
22.3817 | -105.0282
22.3854 | -105.0331
22.3891 | -105.0377
22.3928 | -105.0401
22.3966 | -105.043
22.4003 | -105.0494
22.404 | -105.0544
22.4078 | -105.0597
22.4115 | -105.0653
22.4152 | -105.0699
22.4189 | -105.0754
22.4227 | -105.0786
22.4264 | -105.0793
22.4301 | -105.0811
22.4338 | -105.0834
22.4376 | -105.0846
22.4413 | -105.0859
22.445 | -105.0859
22.4488 | -105.0856
22.4525 | -105.0854
22.4562 | -105.085
22.4599 | -105.0838
22.4637 | -105.0832
22.4674 | -105.0776
22.4711 | -105.0726
22.4748 | -105.0716
22.4786 | -105.0672
22.4823 | -105.0612
22.486 | -105.0555
22.4898 | -105.0486
22.4935 | -105.0351
22.4972 | -105.0209
22.5009 | -105.0029
22.5047 | -104.9829
22.5084 | -104.9507
22.5121 | -104.9123
22.5158 | -104.8721
22.5196 | -104.8345
22.5233 | -104.7948
22.527 | -104.7541
22.5307 | -104.707
22.5345 | -104.6529
22.5382 | -104.6013
22.5419 | -104.5525
22.5457 | -104.4994
22.5494 | -104.4428
22.5531 | -104.3883
22.5568 | -104.3316
22.5606 | -104.2707
22.5643 | -104.2088
22.568 | -104.1466
22.5717 | -104.0823
22.5755 | -104.0156
22.5792 | -103.9467
22.5829 | -103.8619
22.5867 | -103.7936
22.5904 | -103.7238
22.5941 | -103.6546
22.5978 | -103.5866
22.6016 | -103.5175
22.6053 | -103.4497
22.609 | -103.3834
22.6127 | -103.3184
22.6165 | -103.2561
22.6202 | -103.1953
22.6239 | -103.1374
22.6277 | -103.0815
22.6314 | -103.03
22.6351 | -102.983
22.6388 | -102.9389
22.6426 | -102.8971
22.6463 | -102.8584
22.65 | -102.8239
22.6537 | -102.7941
22.6575 | -102.7688
22.6612 | -102.7484
22.6649 | -102.7337
22.6687 | -102.7233
22.6724 | -102.7171
22.6761 | -102.7154
22.6798 | -102.7188
22.6836 | -102.7277
22.6873 | -102.741
22.691 | -102.7598
22.6947 | -102.7806
22.6985 | -102.8071
22.7022 | -102.8393
22.7059 | -102.875
22.7097 | -102.9155
22.7134 | -102.9609
22.7171 | -103.0085
22.7208 | -103.0584
22.7246 | -103.1117
22.7283 | -103.1664
22.732 | -103.2239
22.7357 | -103.2836
22.7395 | -103.3434
22.7432 | -103.4063
22.7469 | -103.4731
22.7507 | -103.5394
22.7544 | -103.6041
22.7581 | -103.6706
22.7618 | -103.7373
22.7656 | -103.8038
22.7693 | -103.8679
22.773 | -103.9318
22.7767 | -103.9948
22.7805 | -104.0559
22.7842 | -104.1138
22.7879 | -104.1718
22.7917 | -104.2271
22.7954 | -104.279
22.7991 | -104.3279
22.8028 | -104.3739
22.8066 | -104.4194
22.8103 | -104.4631
22.814 | -104.5091
22.8177 | -104.5499
22.8215 | -104.5884
22.8252 | -104.6243
22.8289 | -104.6651
22.8327 | -104.704
22.8364 | -104.7373
22.8401 | -104.7626
22.8438 | -104.7855
22.8476 | -104.8024
22.8513 | -104.8196
22.855 | -104.8379
22.8587 | -104.8495
22.8625 | -104.8531
22.8662 | -104.8568
22.8699 | -104.8606
22.8736 | -104.8593
22.8774 | -104.8582
22.8811 | -104.8576
22.8848 | -104.8552
22.8886 | -104.8529
22.8923 | -104.8509
22.896 | -104.8488
22.8997 | -104.849
22.9035 | -104.8478
22.9072 | -104.8455
22.9109 | -104.8427
22.9146 | -104.842
22.9184 | -104.8403
22.9221 | -104.8366
22.9258 | -104.8341
22.9296 | -104.831
22.9333 | -104.8281
22.937 | -104.8248
22.9407 | -104.821
22.9445 | -104.8188
22.9482 | -104.8165
22.9519 | -104.8152
22.9556 | -104.8119
22.9594 | -104.8079
22.9631 | -104.8047
22.9668 | -104.8034
22.9706 | -104.8022
22.9743 | -104.8017
22.978 | -104.8003
22.9817 | -104.7999
22.9855 | -104.7981
22.9892 | -104.795
22.9929 | -104.7937
22.9966 | -104.7935
23.0004 | -104.7918
23.0041 | -104.7898
23.0078 | -104.7889
23.0116 | -104.7885
23.0153 | -104.7847
23.019 | -104.7856
23.0227 | -104.7856
23.0265 | -104.7846
23.0302 | -104.7851
23.0339 | -104.7857
23.0376 | -104.7857
23.0414 | -104.7861
23.0451 | -104.7839
23.0488 | -104.7829
23.0526 | -104.7846
23.0563 | -104.7846
23.06 | -104.7873
23.0637 | -104.7882
23.0675 | -104.7873
23.0712 | -104.783
23.0749 | -104.7792
23.0786 | -104.7792
23.0824 | -104.7763
23.0861 | -104.7715
23.0898 | -104.767
23.0936 | -104.7645
23.0973 | -104.7599
23.101 | -104.7537
23.1047 | -104.7473
23.1085 | -104.742
23.1122 | -104.7369
23.1159 | -104.7317
23.1196 | -104.7238
23.1234 | -104.7157
23.1271 | -104.7109
23.1308 | -104.7044
23.1346 | -104.6959
23.1383 | -104.6889
23.142 | -104.682
23.1457 | -104.6772
23.1495 | -104.6721
23.1532 | -104.6662
23.1569 | -104.6607
23.1606 | -104.6549
23.1644 | -104.6469
23.1681 | -104.6384
23.1718 | -104.6304
23.1756 | -104.6257
23.1793 | -104.6222
23.183 | -104.618
23.1867 | -104.6126
23.1905 | -104.6082
23.1942 | -104.6046
23.1979 | -104.5988
23.2016 | -104.5946
23.2054 | -104.5896
23.2091 | -104.5829
23.2128 | -104.5772
23.2165 | -104.5734
23.2203 | -104.5682
23.224 | -104.5634
23.2277 | -104.5593
23.2315 | -104.5558
23.2352 | -104.5514
23.2389 | -104.547
23.2426 | -104.5416
23.2464 | -104.5362
23.2501 | -104.5316
23.2538 | -104.5267
23.2575 | -104.5208
23.2613 | -104.5137
23.265 | -104.5059
23.2687 | -104.4965
23.2725 | -104.4882
23.2762 | -104.4811
23.2799 | -104.4753
23.2836 | -104.4701
23.2874 | -104.4641
23.2911 | -104.4559
23.2948 | -104.4484
23.2985 | -104.441
23.3023 | -104.4323
23.306 | -104.4234
23.3097 | -104.4154
23.3135 | -104.4085
23.3172 | -104.4021
23.3209 | -104.3966
23.3246 | -104.3925
23.3284 | -104.3903
23.3321 | -104.3883
23.3358 | -104.3847
23.3395 | -104.3837
23.3433 | -104.3837
23.347 | -104.3814
23.3507 | -104.3795
23.3545 | -104.3778
23.3582 | -104.3758
23.3619 | -104.3737
23.3656 | -104.3728
23.3694 | -104.3736
23.3731 | -104.3773
23.3768 | -104.38
23.3805 | -104.3837
23.3843 | -104.3889
23.388 | -104.3923
23.3917 | -104.3979
23.3955 | -104.4035
23.3992 | -104.4106
23.4029 | -104.4203
23.4066 | -104.4281
23.4104 | -104.4344
23.4141 | -104.4425
23.4178 | -104.4494
23.4215 | -104.4564
23.4253 | -104.4635
23.429 | -104.4716
23.4327 | -104.48
23.4365 | -104.4893
23.4402 | -104.4963
23.4439 | -104.5025
23.4476 | -104.5104
23.4514 | -104.5187
23.4551 | -104.5261
23.4588 | -104.5336
23.4625 | -104.5438
23.4663 | -104.5538
23.47 | -104.5625
23.4737 | -104.5704
23.4775 | -104.5778
23.4812 | -104.5868
23.4849 | -104.5955
23.4886 | -104.6045
23.4924 | -104.6141
23.4961 | -104.6252
23.4998 | -104.6371
23.5035 | -104.6481
23.5073 | -104.6571
23.511 | -104.6655
23.5147 | -104.6745
23.5184 | -104.6838
23.5222 | -104.693
23.5259 | -104.7027
23.5296 | -104.7152
23.5334 | -104.7272
23.5371 | -104.7387
23.5408 | -104.7524
23.5445 | -104.7664
23.5483 | -104.7776
23.552 | -104.7907
23.5557 | -104.8018
23.5594 | -104.8114
23.5632 | -104.8225
23.5669 | -104.8347
23.5706 | -104.8471
23.5744 | -104.8593
23.5781 | -104.8714
23.5818 | -104.8838
23.5855 | -104.897
23.5893 | -104.9094
23.593 | -104.9216
23.5967 | -104.9361
23.6004 | -104.952
23.6042 | -104.9666
23.6079 | -104.9802
23.6116 | -104.9918
23.6154 | -105.0045
23.6191 | -105.0189
23.6228 | -105.0355
23.6265 | -105.0508
23.6303 | -105.0672
23.634 | -105.0829
23.6377 | -105.098
23.6414 | -105.1127
23.6452 | -105.1268
23.6489 | -105.1411
23.6526 | -105.1547
23.6564 | -105.1663
23.6601 | -105.1774
23.6638 | -105.187
23.6675 | -105.1976
23.6713 | -105.2089
23.675 | -105.2187
23.6787 | -105.2287
23.6824 | -105.2394
23.6862 | -105.2502
23.6899 | -105.2615
23.6936 | -105.2726
23.6974 | -105.2823
23.7011 | -105.2897
23.7048 | -105.2976
23.7085 | -105.3038
23.7123 | -105.3091
23.716 | -105.3122
23.7197 | -105.3155
23.7234 | -105.3186
23.7272 | -105.3211
23.7309 | -105.3205
23.7346 | -105.321
23.7384 | -105.3179
23.7421 | -105.3166
23.7458 | -105.3156
23.7495 | -105.3142
23.7533 | -105.3136
23.757 | -105.3126
23.7607 | -105.3103
23.7644 | -105.3021
23.7682 | -105.2933
23.7719 | -105.2849
23.7756 | -105.2783
23.7794 | -105.2705
23.7831 | -105.2596
23.7868 | -105.2484
23.7905 | -105.2373
23.7943 | -105.2242
23.798 | -105.2104
23.8017 | -105.1958
23.8054 | -105.1803
23.8092 | -105.1645
23.8129 | -105.149
23.8166 | -105.1332
23.8204 | -105.1164
23.8241 | -105.0968
23.8278 | -105.0757
23.8315 | -105.056
23.8353 | -105.0351
23.839 | -105.0151
23.8427 | -104.9948
23.8464 | -104.9741
23.8502 | -104.9554
23.8539 | -104.937
23.8576 | -104.9187
23.8613 | -104.9003
23.8651 | -104.8806
23.8688 | -104.8615
23.8725 | -104.8453
23.8763 | -104.829
23.88 | -104.8115
23.8837 | -104.796
23.8874 | -104.7812
23.8912 | -104.7687
23.8949 | -104.7575
23.8986 | -104.747
23.9023 | -104.7368
23.9061 | -104.7269
23.9098 | -104.7192
23.9135 | -104.7109
23.9173 | -104.7044
23.921 | -104.7005
23.9247 | -104.6979
23.9284 | -104.6972
23.9322 | -104.6953
23.9359 | -104.6943
23.9396 | -104.695
23.9433 | -104.6978
23.9471 | -104.7022
23.9508 | -104.7072
23.9545 | -104.7138
23.9583 | -104.7199
23.962 | -104.7265
23.9657 | -104.734
23.9694 | -104.7435
23.9732 | -104.754
23.9769 | -104.7638
23.9806 | -104.7748
23.9843 | -104.7859
23.9881 | -104.7967
23.9918 | -104.808
23.9955 | -104.8207
23.9993 | -104.8342
24.003 | -104.8468
24.0067 | -104.8585
24.0104 | -104.8707
24.0142 | -104.8819
24.0179 | -104.8927
24.0216 | -104.9029
24.0253 | -104.9126
24.0291 | -104.9233
24.0328 | -104.9333
24.0365 | -104.9439
24.0403 | -104.9551
24.044 | -104.9634
24.0477 | -104.9704
24.0514 | -104.9791
24.0552 | -104.9871
24.0589 | -104.9946
24.0626 | -104.9993
24.0663 | -105.0028
24.0701 | -105.0097
24.0738 | -105.0145
24.0775 | -105.0184
24.0813 | -105.0202
24.085 | -105.0229
24.0887 | -105.0242
24.0924 | -105.0243
24.0962 | -105.0228
24.0999 | -105.0221
24.1036 | -105.0254
24.1073 | -105.0267
24.1111 | -105.0258
24.1148 | -105.0249
24.1185 | -105.0246
24.1223 | -105.0232
24.126 | -105.0216
24.1297 | -105.0203
24.1334 | -105.0176
24.1372 | -105.0133
24.1409 | -105.0107
24.1446 | -105.0096
24.1483 | -105.008
24.1521 | -105.0065
24.1558 | -105.0055
24.1595 | -105.0049
24.1633 | -105.0064
24.167 | -105.0078
24.1707 | -105.0076
24.1744 | -105.007
24.1782 | -105.0059
24.1819 | -105.0055
24.1856 | -105.0062
24.1893 | -105.0055
24.1931 | -105.005
24.1968 | -105.0087
24.2005 | -105.0078
24.2042 | -105.0083
24.208 | -105.011
24.2117 | -105.0139
24.2154 | -105.0153
24.2192 | -105.0168
24.2229 | -105.0199
24.2266 | -105.0214
24.2303 | -105.0228
24.2341 | -105.0232
24.2378 | -105.024
24.2415 | -105.0245
24.2452 | -105.0256
24.249 | -105.0276
24.2527 | -105.0298
24.2564 | -105.0326
24.2602 | -105.0308
24.2639 | -105.0283
24.2676 | -105.0267
24.2713 | -105.0231
24.2751 | -105.0181
24.2788 | -105.0085
24.2825 | -104.9942
24.2862 | -104.9777
24.29 | -104.9573
24.2937 | -104.9345
24.2974 | -104.9107
24.3012 | -104.8895
24.3049 | -104.8712
24.3086 | -104.8526
24.3123 | -104.833
24.3161 | -104.8106
24.3198 | -104.7878
24.3235 | -104.763
24.3272 | -104.7366
24.331 | -104.7088
24.3347 | -104.6799
24.3384 | -104.6508
24.3422 | -104.6213
24.3459 | -104.5913
24.3496 | -104.5591
24.3533 | -104.5255
24.3571 | -104.4931
24.3608 | -104.4608
24.3645 | -104.428
24.3682 | -104.3933
24.372 | -104.3582
24.3757 | -104.3242
24.3794 | -104.2899
24.3832 | -104.2546
24.3869 | -104.2196
24.3906 | -104.1857
24.3943 | -104.1496
24.3981 | -104.1146
24.4018 | -104.0823
24.4055 | -104.0527
24.4092 | -104.0255
24.413 | -103.9981
24.4167 | -103.972
24.4204 | -103.9481
24.4242 | -103.9258
24.4279 | -103.9045
24.4316 | -103.8842
24.4353 | -103.8663
24.4391 | -103.8514
24.4428 | -103.8392
24.4465 | -103.8288
24.4502 | -103.8209
24.454 | -103.8153
24.4577 | -103.8123
24.4614 | -103.8113
24.4652 | -103.8141
24.4689 | -103.8189
24.4726 | -103.8264
24.4763 | -103.8356
24.4801 | -103.8474
24.4838 | -103.8586
24.4875 | -103.8703
24.4912 | -103.8857
24.495 | -103.9033
24.4987 | -103.922
24.5024 | -103.9427
24.5062 | -103.9647
24.5099 | -103.9869
24.5136 | -104.0081
24.5173 | -104.0311
24.5211 | -104.056
24.5248 | -104.0812
24.5285 | -104.1069
24.5322 | -104.1322
24.536 | -104.1589
24.5397 | -104.1837
24.5434 | -104.2085
24.5471 | -104.2337
24.5509 | -104.2575
24.5546 | -104.2793
24.5583 | -104.3003
24.5621 | -104.3192
24.5658 | -104.3347
24.5695 | -104.3447
24.5732 | -104.348
24.577 | -104.3305
24.5807 | -104.2915
24.5844 | -104.2319
24.5881 | -104.1468
24.5919 | -104.0666
24.5956 | -104.0011
24.5993 | -103.9314
24.6031 | -103.8706
24.6068 | -103.81
24.6105 | -103.7386
24.6142 | -103.6581
24.618 | -103.5701
24.6217 | -103.4766
24.6254 | -103.3831
24.6291 | -103.2883
24.6329 | -103.1914
24.6366 | -103.0932
24.6403 | -102.9894
24.6441 | -102.8814
24.6478 | -102.7685
24.6515 | -102.6505
24.6552 | -102.5299
24.659 | -102.4082
24.6627 | -102.2865
24.6664 | -102.1644
24.6701 | -102.0429
24.6739 | -101.9223
24.6776 | -101.8032
24.6813 | -101.6856
24.6851 | -101.5692
24.6888 | -101.4553
24.6925 | -101.3454
24.6962 | -101.2391
24.7 | -101.1365
24.7037 | -101.0377
24.7074 | -100.9436
24.7111 | -100.8543
24.7149 | -100.7707
24.7186 | -100.6922
24.7223 | -100.6193
24.7261 | -100.5522
24.7298 | -100.4916
24.7335 | -100.4841
24.7372 | -100.4384
24.741 | -100.3988
24.7447 | -100.3659
24.7484 | -100.3383
24.7521 | -100.316
24.7559 | -100.2999
24.7596 | -100.291
24.7633 | -100.2891
24.7671 | -100.2933
24.7708 | -100.3041
24.7745 | -100.3213
24.7782 | -100.3457
24.782 | -100.377
24.7857 | -100.4141
24.7894 | -100.4575
24.7931 | -100.507
24.7969 | -100.5622
24.8006 | -100.6227
24.8043 | -100.6893
24.8081 | -100.7614
24.8118 | -100.8387
24.8155 | -100.9202
24.8192 | -101.0065
24.823 | -101.0972
24.8267 | -101.1922
24.8304 | -101.2907
24.8341 | -101.394
24.8379 | -101.5011
24.8416 | -101.6097
24.8453 | -101.7193
24.849 | -101.8308
24.8528 | -101.9439
24.8565 | -102.0573
24.8602 | -102.1722
24.864 | -102.2889
24.8677 | -102.4038
24.8714 | -102.5163
24.8751 | -102.628
24.8789 | -102.7368
24.8826 | -102.8454
24.8863 | -102.9505
24.89 | -103.0521
24.8938 | -103.1516
24.8975 | -103.2485
24.9012 | -103.3411
24.905 | -103.4288
24.9087 | -103.5116
24.9124 | -103.5907
24.9161 | -103.6679
24.9199 | -103.7426
24.9236 | -103.8172
24.9273 | -103.899
24.931 | -103.9879
24.9348 | -104.0803
24.9385 | -104.1782
24.9422 | -104.2542
24.946 | -104.2991
24.9497 | -104.336
24.9534 | -104.355
24.9571 | -104.3673
24.9609 | -104.3815
24.9646 | -104.3977
24.9683 | -104.4155
24.972 | -104.4307
24.9758 | -104.4391
24.9795 | -104.4439
24.9832 | -104.4457
24.987 | -104.4459
24.9907 | -104.4463
24.9944 | -104.4465
24.9981 | -104.4479
25.0019 | -104.4513
25.0056 | -104.4534
25.0093 | -104.4564
25.013 | -104.4612
25.0168 | -104.4661
25.0205 | -104.4716
25.0242 | -104.4776
25.028 | -104.4835
25.0317 | -104.4908
25.0354 | -104.4994
25.0391 | -104.5099
25.0429 | -104.5203
25.0466 | -104.5304
25.0503 | -104.5406
25.054 | -104.5513
25.0578 | -104.5623
25.0615 | -104.5735
25.0652 | -104.585
25.069 | -104.5978
25.0727 | -104.6112
25.0764 | -104.6253
25.0801 | -104.6391
25.0839 | -104.6523
25.0876 | -104.6661
25.0913 | -104.6804
25.095 | -104.6953
25.0988 | -104.7113
25.1025 | -104.7289
25.1062 | -104.7464
25.11 | -104.7647
25.1137 | -104.7815
25.1174 | -104.7958
25.1211 | -104.8093
25.1249 | -104.8218
25.1286 | -104.8348
25.1323 | -104.8454
25.136 | -104.855
25.1398 | -104.865
25.1435 | -104.8744
25.1472 | -104.8829
25.151 | -104.8908
25.1547 | -104.9008
25.1584 | -104.91
25.1621 | -104.9163
25.1659 | -104.9223
25.1696 | -104.9285
25.1733 | -104.9327
25.177 | -104.9364
25.1808 | -104.9378
25.1845 | -104.9392
25.1882 | -104.9404
25.1919 | -104.9402
25.1957 | -104.9404
25.1994 | -104.9409
25.2031 | -104.9378
25.2069 | -104.9322
25.2106 | -104.9272
25.2143 | -104.9224
25.218 | -104.9155
25.2218 | -104.9097
25.2255 | -104.902
25.2292 | -104.8916
25.2329 | -104.8823
25.2367 | -104.8697
25.2404 | -104.8567
25.2441 | -104.8465
25.2479 | -104.8328
25.2516 | -104.8177
25.2553 | -104.8044
25.259 | -104.7918
25.2628 | -104.7807
25.2665 | -104.7674
25.2702 | -104.7459
25.2739 | -104.7317
25.2777 | -104.7177
25.2814 | -104.7021
25.2851 | -104.6882
25.2889 | -104.6747
25.2926 | -104.6593
25.2963 | -104.6449
25.3 | -104.6299
25.3038 | -104.6168
25.3075 | -104.6046
25.3112 | -104.5913
25.3149 | -104.5782
25.3187 | -104.5659
25.3224 | -104.5544
25.3261 | -104.5432
25.3299 | -104.5337
25.3336 | -104.5244
25.3373 | -104.516
25.341 | -104.507
25.3448 | -104.4996
25.3485 | -104.4936
25.3522 | -104.4875
25.3559 | -104.4807
25.3597 | -104.4746
25.3634 | -104.4718
25.3671 | -104.4694
25.3709 | -104.4676
25.3746 | -104.4671
25.3783 | -104.4671
25.382 | -104.4673
25.3858 | -104.4679
25.3895 | -104.4684
25.3932 | -104.4699
25.3969 | -104.4708
25.4007 | -104.4723
25.4044 | -104.4735
25.4081 | -104.4756
25.4119 | -104.4797
25.4156 | -104.4852
25.4193 | -104.4893
25.423 | -104.4927
25.4268 | -104.4961
25.4305 | -104.5006
25.4342 | -104.5053
25.4379 | -104.5106
25.4417 | -104.5151
25.4454 | -104.5183
25.4491 | -104.5206
25.4529 | -104.5231
25.4566 | -104.5264
25.4603 | -104.53
25.464 | -104.5344
25.4678 | -104.5364
25.4715 | -104.5393
25.4752 | -104.5427
25.4789 | -104.5461
25.4827 | -104.5511
25.4864 | -104.5541
25.4901 | -104.5568
25.4938 | -104.5602
25.4976 | -104.5637
25.5013 | -104.5652
25.505 | -104.5661
25.5088 | -104.5674
25.5125 | -104.5672
25.5162 | -104.5666
25.5199 | -104.5664
25.5237 | -104.5659
25.5274 | -104.5654
25.5311 | -104.5658
25.5348 | -104.5641
25.5386 | -104.5622
25.5423 | -104.5598
25.546 | -104.5564
25.5498 | -104.5532
25.5535 | -104.5491
25.5572 | -104.5443
25.5609 | -104.5411
25.5647 | -104.5393
25.5684 | -104.5363
25.5721 | -104.5301
25.5758 | -104.5249
25.5796 | -104.5197
25.5833 | -104.5145
25.587 | -104.5094
25.5908 | -104.5057
25.5945 | -104.5018
25.5982 | -104.4973
25.6019 | -104.4927
25.6057 | -104.488
25.6094 | -104.4837
25.6131 | -104.4776
25.6168 | -104.4717
25.6206 | -104.4653
25.6243 | -104.4593
25.628 | -104.4561
25.6318 | -104.4548
25.6355 | -104.4505
25.6392 | -104.4461
25.6429 | -104.4425
25.6467 | -104.4391
25.6504 | -104.4355
25.6541 | -104.4321
25.6578 | -104.4281
25.6616 | -104.4243
25.6653 | -104.4205
25.669 | -104.4165
25.6728 | -104.4114
25.6765 | -104.4091
25.6802 | -104.408
25.6839 | -104.4069
25.6877 | -104.405
25.6914 | -104.4031
25.6951 | -104.4029
25.6988 | -104.4013
25.7026 | -104.4014
25.7063 | -104.4022
25.71 | -104.4014
25.7138 | -104.4001
25.7175 | -104.401
25.7212 | -104.4012
25.7249 | -104.4016
25.7287 | -104.403
25.7324 | -104.4031
25.7361 | -104.4031
25.7398 | -104.4043
25.7436 | -104.4058
25.7473 | -104.4075
25.751 | -104.4095
25.7548 | -104.4112
25.7585 | -104.414
25.7622 | -104.4175
25.7659 | -104.4227
25.7697 | -104.4271
25.7734 | -104.4297
25.7771 | -104.431
25.7808 | -104.4332
25.7846 | -104.4369
25.7883 | -104.4406
25.792 | -104.4438
25.7958 | -104.4479
25.7995 | -104.4512
25.8032 | -104.455
25.8069 | -104.4555
25.8107 | -104.4595
25.8144 | -104.4645
25.8181 | -104.4675
25.8218 | -104.4694
25.8256 | -104.4716
25.8293 | -104.4743
25.833 | -104.4782
25.8367 | -104.4803
25.8405 | -104.4823
25.8442 | -104.4847
25.8479 | -104.4857
25.8517 | -104.4872
25.8554 | -104.4906
25.8591 | -104.4933
25.8628 | -104.495
25.8666 | -104.4974
25.8703 | -104.4987
25.874 | -104.4997
25.8777 | -104.5022
25.8815 | -104.5032
25.8852 | -104.5047
25.8889 | -104.5068
25.8927 | -104.5094
25.8964 | -104.5112
25.9001 | -104.5136
25.9038 | -104.5174
25.9076 | -104.5205
25.9113 | -104.5225
25.915 | -104.5239
25.9187 | -104.5252
25.9225 | -104.5266
25.9262 | -104.5281
25.9299 | -104.5282
25.9337 | -104.5304
25.9374 | -104.5326
25.9411 | -104.5337
25.9448 | -104.5359
25.9486 | -104.5389
25.9523 | -104.541
25.956 | -104.5403
25.9597 | -104.5406
25.9635 | -104.5429
25.9672 | -104.5452
25.9709 | -104.5491
25.9747 | -104.5519
25.9784 | -104.5557
25.9821 | -104.5587
25.9858 | -104.5626
25.9896 | -104.5677
25.9933 | -104.572
25.997 | -104.5735
26.0007 | -104.5764
26.0045 | -104.5803
26.0082 | -104.5849
26.0119 | -104.5885
26.0157 | -104.5918
26.0194 | -104.5962
26.0231 | -104.601
26.0268 | -104.6051
26.0306 | -104.609
26.0343 | -104.6121
26.038 | -104.618
26.0417 | -104.6216
26.0455 | -104.6248
26.0492 | -104.6283
26.0529 | -104.632
26.0567 | -104.6347
26.0604 | -104.6364
26.0641 | -104.6402
26.0678 | -104.642
26.0716 | -104.6427
26.0753 | -104.6454
26.079 | -104.6486
26.0827 | -104.6496
26.0865 | -104.651
26.0902 | -104.6513
26.0939 | -104.6516
26.0977 | -104.6523
26.1014 | -104.6529
26.1051 | -104.6519
26.1088 | -104.6507
26.1126 | -104.6503
26.1163 | -104.6484
26.12 | -104.6472
26.1237 | -104.6433
26.1275 | -104.6404
26.1312 | -104.6385
26.1349 | -104.6355
26.1387 | -104.6314
26.1424 | -104.6279
26.1461 | -104.6256
26.1498 | -104.6211
26.1536 | -104.6167
26.1573 | -104.611
26.161 | -104.6039
26.1647 | -104.5973
26.1685 | -104.5914
26.1722 | -104.5871
26.1759 | -104.5825
26.1796 | -104.5774
26.1834 | -104.5699
26.1871 | -104.5612
26.1908 | -104.5535
26.1946 | -104.5456
26.1983 | -104.5367
26.202 | -104.5285
26.2057 | -104.5197
26.2095 | -104.5119
26.2132 | -104.504
26.2169 | -104.4959
26.2206 | -104.4888
26.2244 | -104.4805
26.2281 | -104.4732
26.2318 | -104.4662
26.2356 | -104.4576
26.2393 | -104.451
26.243 | -104.4446
26.2467 | -104.4363
26.2505 | -104.4294
26.2542 | -104.4241
26.2579 | -104.4181
26.2616 | -104.4122
26.2654 | -104.4058
26.2691 | -104.4006
26.2728 | -104.3954
26.2766 | -104.3903
26.2803 | -104.3845
26.284 | -104.3799
26.2877 | -104.3769
26.2915 | -104.3745
26.2952 | -104.3731
26.2989 | -104.3708
26.3026 | -104.3661
26.3064 | -104.3633
26.3101 | -104.3608
26.3138 | -104.3586
26.3176 | -104.3562
26.3213 | -104.3548
26.325 | -104.356
26.3287 | -104.3577
26.3325 | -104.3569
26.3362 | -104.3541
26.3399 | -104.3538
26.3436 | -104.3553
26.3474 | -104.3622
26.3511 | -104.3647
26.3548 | -104.3677
26.3586 | -104.3707
26.3623 | -104.3744
26.366 | -104.3782
26.3697 | -104.3816
26.3735 | -104.3844
26.3772 | -104.3873
26.3809 | -104.3901
26.3846 | -104.3936
26.3884 | -104.399
26.3921 | -104.4038
26.3958 | -104.4087
26.3996 | -104.4142
26.4033 | -104.4194
26.407 | -104.4238
26.4107 | -104.4256
26.4145 | -104.4277
26.4182 | -104.4296
26.4219 | -104.4306
26.4256 | -104.433
26.4294 | -104.4359
26.4331 | -104.4385
26.4368 | -104.4395
26.4406 | -104.4401
26.4443 | -104.4399
26.448 | -104.4393
26.4517 | -104.4389
26.4555 | -104.439
26.4592 | -104.4389
26.4629 | -104.4367
26.4666 | -104.4331
26.4704 | -104.429
26.4741 | -104.4253
26.4778 | -104.4225
26.4816 | -104.4189
26.4853 | -104.4156
26.489 | -104.4142
26.4927 | -104.4103
26.4965 | -104.4061
26.5002 | -104.4007
26.5039 | -104.3942
26.5076 | -104.3889
26.5114 | -104.3835
26.5151 | -104.3795
26.5188 | -104.3755
26.5225 | -104.3709
26.5263 | -104.3659
26.53 | -104.3627
26.5337 | -104.3601
26.5375 | -104.3572
26.5412 | -104.3559
26.5449 | -104.354
26.5486 | -104.3507
26.5524 | -104.3461
26.5561 | -104.3448
26.5598 | -104.3459
26.5635 | -104.3464
26.5673 | -104.3459
26.571 | -104.3461
26.5747 | -104.3462
26.5785 | -104.347
26.5822 | -104.3481
26.5859 | -104.3504
26.5896 | -104.3524
26.5934 | -104.3532
26.5971 | -104.3552
26.6008 | -104.3564
26.6045 | -104.3582
26.6083 | -104.3619
26.612 | -104.3646
26.6157 | -104.3682
26.6195 | -104.3731
26.6232 | -104.3777
26.6269 | -104.3812
26.6306 | -104.3847
26.6344 | -104.3882
26.6381 | -104.3944
26.6418 | -104.4013
26.6455 | -104.4087
26.6493 | -104.4154
26.653 | -104.4226
26.6567 | -104.4294
26.6605 | -104.4362
26.6642 | -104.4429
26.6679 | -104.4484
26.6716 | -104.457
26.6754 | -104.4661
26.6791 | -104.4741
26.6828 | -104.4834
26.6865 | -104.4905
26.6903 | -104.4983
26.694 | -104.5062
26.6977 | -104.5129
26.7015 | -104.5187
26.7052 | -104.5259
26.7089 | -104.535
26.7126 | -104.5405
26.7164 | -104.5453
26.7201 | -104.55
26.7238 | -104.5547
26.7275 | -104.5603
26.7313 | -104.5659
26.735 | -104.572
26.7387 | -104.5768
26.7425 | -104.581
26.7462 | -104.5876
26.7499 | -104.5944
26.7536 | -104.5999
26.7574 | -104.6041
26.7611 | -104.6088
26.7648 | -104.6129
26.7685 | -104.6179
26.7723 | -104.6237
26.776 | -104.6276
26.7797 | -104.6303
26.7835 | -104.6356
26.7872 | -104.6421
26.7909 | -104.6475
26.7946 | -104.6528
26.7984 | -104.6581
26.8021 | -104.663
26.8058 | -104.6685
26.8095 | -104.6728
26.8133 | -104.6777
26.817 | -104.6832
26.8207 | -104.6882
26.8244 | -104.6951
26.8282 | -104.7005
26.8319 | -104.7046
26.8356 | -104.7086
26.8394 | -104.7159
26.8431 | -104.7227
26.8468 | -104.7294
26.8505 | -104.7354
26.8543 | -104.738
26.858 | -104.7402
26.8617 | -104.7429
26.8654 | -104.745
26.8692 | -104.7468
26.8729 | -104.7503
26.8766 | -104.7527
26.8804 | -104.7547
26.8841 | -104.749
26.8878 | -104.7495
26.8915 | -104.7491
26.8953 | -104.7459
26.899 | -104.7422
26.9027 | -104.7388
26.9064 | -104.7352
26.9102 | -104.7296
26.9139 | -104.7253
26.9176 | -104.7209
26.9214 | -104.7135
26.9251 | -104.7058
26.9288 | -104.7006
26.9325 | -104.6958
26.9363 | -104.6888
26.94 | -104.6804
26.9437 | -104.6714
26.9474 | -104.6628
26.9512 | -104.6538
26.9549 | -104.6439
26.9586 | -104.633
26.9624 | -104.6205
26.9661 | -104.6083
26.9698 | -104.5979
26.9735 | -104.5878
26.9773 | -104.5762
26.981 | -104.5648
26.9847 | -104.554
26.9884 | -104.5418
26.9922 | -104.5301
26.9959 | -104.5186
26.9996 | -104.5067
27.0034 | -104.4934
27.0071 | -104.4812
27.0108 | -104.4707
27.0145 | -104.4603
27.0183 | -104.4495
27.022 | -104.4359
27.0257 | -104.4223
27.0294 | -104.4107
27.0332 | -104.3992
27.0369 | -104.3875
27.0406 | -104.3766
27.0444 | -104.3677
27.0481 | -104.3592
27.0518 | -104.3497
27.0555 | -104.3388
27.0593 | -104.3292
27.063 | -104.3212
27.0667 | -104.3159
27.0704 | -104.311
27.0742 | -104.3052
27.0779 | -104.2985
27.0816 | -104.2924
27.0854 | -104.2888
27.0891 | -104.2841
27.0928 | -104.281
27.0965 | -104.2784
27.1003 | -104.2763
27.104 | -104.2722
27.1077 | -104.2688
27.1114 | -104.2679
27.1152 | -104.2675
27.1189 | -104.2683
27.1226 | -104.2704
27.1264 | -104.27
27.1301 | -104.2684
27.1338 | -104.2668
27.1375 | -104.2665
27.1413 | -104.2667
27.145 | -104.2663
27.1487 | -104.2653
27.1524 | -104.2635
27.1562 | -104.263
27.1599 | -104.2655
27.1636 | -104.2656
27.1673 | -104.2656
27.1711 | -104.267
27.1748 | -104.2667
27.1785 | -104.2672
27.1823 | -104.2671
27.186 | -104.2663
27.1897 | -104.2659
27.1934 | -104.2646
27.1972 | -104.2628
27.2009 | -104.262
27.2046 | -104.2594
27.2083 | -104.2578
27.2121 | -104.2555
27.2158 | -104.2535
27.2195 | -104.2491
27.2233 | -104.2464
27.227 | -104.2446
27.2307 | -104.2429
27.2344 | -104.2436
27.2382 | -104.2438
27.2419 | -104.2416
27.2456 | -104.2393
27.2493 | -104.2376
27.2531 | -104.2362
27.2568 | -104.2356
27.2605 | -104.2347
27.2643 | -104.2342
27.268 | -104.2331
27.2717 | -104.232
27.2754 | -104.231
27.2792 | -104.2288
27.2829 | -104.2266
27.2866 | -104.2257
27.2903 | -104.2243
27.2941 | -104.2233
27.2978 | -104.2213
27.3015 | -104.2187
27.3053 | -104.2171
27.309 | -104.2165
27.3127 | -104.2152
27.3164 | -104.2141
27.3202 | -104.2138
27.3239 | -104.2146
27.3276 | -104.2163
27.3313 | -104.217
27.3351 | -104.2165
27.3388 | -104.2169
27.3425 | -104.2199
27.3463 | -104.2237
27.35 | -104.2261
27.3537 | -104.2258
27.3574 | -104.2274
27.3612 | -104.2303
27.3649 | -104.2339
27.3686 | -104.2365
27.3723 | -104.2393
27.3761 | -104.2422
27.3798 | -104.2442
27.3835 | -104.2476
27.3873 | -104.2526
27.391 | -104.2563
27.3947 | -104.2595
27.3984 | -104.2617
27.4022 | -104.2636
27.4059 | -104.2682
27.4096 | -104.2735
27.4133 | -104.2759
27.4171 | -104.2789
27.4208 | -104.2801
27.4245 | -104.2832
27.4283 | -104.2847
27.432 | -104.2853
27.4357 | -104.2865
27.4394 | -104.288
27.4432 | -104.2905
27.4469 | -104.2915
27.4506 | -104.2908
27.4543 | -104.2891
27.4581 | -104.2882
27.4618 | -104.2887
27.4655 | -104.2894
27.4693 | -104.2883
27.473 | -104.2855
27.4767 | -104.2845
27.4804 | -104.282
27.4842 | -104.2776
27.4879 | -104.2729
27.4916 | -104.27
27.4953 | -104.2665
27.4991 | -104.2622
27.5028 | -104.2574
27.5065 | -104.2526
27.5102 | -104.2473
27.514 | -104.2432
27.5177 | -104.2399
27.5214 | -104.2366
27.5252 | -104.2334
27.5289 | -104.2288
27.5326 | -104.2241
27.5363 | -104.2204
27.5401 | -104.2169
27.5438 | -104.2125
27.5475 | -104.2072
27.5512 | -104.2035
27.555 | -104.1989
27.5587 | -104.192
27.5624 | -104.1869
27.5662 | -104.1818
27.5699 | -104.1781
27.5736 | -104.1745
27.5773 | -104.1716
27.5811 | -104.1703
27.5848 | -104.1691
27.5885 | -104.1664
27.5922 | -104.1652
27.596 | -104.1641
27.5997 | -104.1634
27.6034 | -104.1637
27.6072 | -104.1652
27.6109 | -104.1658
27.6146 | -104.1665
27.6183 | -104.1663
27.6221 | -104.1651
27.6258 | -104.1663
27.6295 | -104.168
27.6332 | -104.1691
27.637 | -104.1723
27.6407 | -104.1751
27.6444 | -104.1778
27.6482 | -104.1797
27.6519 | -104.1813
27.6556 | -104.185
27.6593 | -104.1886
27.6631 | -104.1928
27.6668 | -104.198
27.6705 | -104.2033
27.6742 | -104.2077
27.678 | -104.2124
27.6817 | -104.2173
27.6854 | -104.2222
27.6892 | -104.2278
27.6929 | -104.2328
27.6966 | -104.2361
27.7003 | -104.2407
27.7041 | -104.2465
27.7078 | -104.2519
27.7115 | -104.2574
27.7152 | -104.2627
27.719 | -104.2692
27.7227 | -104.2764
27.7264 | -104.2813
27.7302 | -104.2839
27.7339 | -104.2865
27.7376 | -104.2905
27.7413 | -104.296
27.7451 | -104.3006
27.7488 | -104.304
27.7525 | -104.3078
27.7562 | -104.3126
27.76 | -104.3163
27.7637 | -104.3189
27.7674 | -104.3226
27.7712 | -104.3261
27.7749 | -104.3287
27.7786 | -104.3301
27.7823 | -104.3299
27.7861 | -104.3319
27.7898 | -104.3347
27.7935 | -104.3363
27.7972 | -104.3379
27.801 | -104.3404
27.8047 | -104.342
27.8084 | -104.3406
27.8122 | -104.3375
27.8159 | -104.3376
27.8196 | -104.3369
27.8233 | -104.3355
27.8271 | -104.3346
27.8308 | -104.3331
27.8345 | -104.3308
27.8382 | -104.3288
27.842 | -104.3268
27.8457 | -104.3246
27.8494 | -104.3219
27.8531 | -104.3195
27.8569 | -104.3173
27.8606 | -104.3154
27.8643 | -104.3129
27.8681 | -104.3094
27.8718 | -104.3074
27.8755 | -104.3062
27.8792 | -104.3036
27.883 | -104.3001
27.8867 | -104.2972
27.8904 | -104.2936
27.8941 | -104.2886
27.8979 | -104.2845
27.9016 | -104.2816
27.9053 | -104.2796
27.9091 | -104.2779
27.9128 | -104.2763
27.9165 | -104.2733
27.9202 | -104.2722
27.924 | -104.2713
27.9277 | -104.2692
27.9314 | -104.2664
27.9351 | -104.2645
27.9389 | -104.2639
27.9426 | -104.2629
27.9463 | -104.2633
27.9501 | -104.2646
27.9538 | -104.2657
27.9575 | -104.2703
27.9612 | -104.2715
27.965 | -104.2717
27.9687 | -104.2731
27.9724 | -104.2733
27.9761 | -104.2726
27.9799 | -104.2732
27.9836 | -104.2757
27.9873 | -104.2778
27.9911 | -104.2792
27.9948 | -104.2813
27.9985 | -104.2817
28.0022 | -104.2839