## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 235 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 237
Hood | 02-92B/P | 176
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1011.9 | hPa |  | PRC PRO 01
Humidity | 47.6 | %/Hr
Internal temperature of calibration bench | 23.18 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1024 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.6 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.91 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.03 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3063 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 23/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 27/05/2018
Engineer | Anthony HUE | Approval |  | 29/05/2018
CEO | Robert GLEMOT | Approval |  | 29/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4623 | 57.5815
0.012209 | -36.0377 | 38.4053
0.018313 | -34.9781 | 28.2553
0.024417 | -34.5134 | 21.9946
0.030521 | -34.3979 | 18.2177
0.036626 | -34.2526 | 15.1665
0.04273 | -34.1561 | 12.8141
0.048834 | -34.1098 | 11.2915
0.054938 | -34.0465 | 10.1527
0.061043 | -33.987 | 9.0763
0.067147 | -33.9761 | 8.2536
0.073251 | -33.9759 | 7.4942
0.079355 | -33.9923 | 6.8341
0.08546 | -34.0046 | 6.2435
0.091564 | -33.9993 | 5.8032
0.097668 | -33.9815 | 5.5488
0.10377 | -33.9602 | 5.2997
0.10988 | -33.9435 | 5.0226
0.11598 | -33.9345 | 4.7317
0.12209 | -33.936 | 4.4493
0.12819 | -33.9452 | 4.1673
0.13429 | -33.9553 | 3.8816
0.1404 | -33.9583 | 3.6079
0.1465 | -33.96 | 3.354
0.15261 | -33.9573 | 3.1371
0.15871 | -33.9511 | 2.976
0.16482 | -33.9433 | 2.8716
0.17092 | -33.9409 | 2.8275
0.17702 | -33.9386 | 2.7521
0.18313 | -33.9375 | 2.6715
0.18923 | -33.937 | 2.5622
0.19534 | -33.9356 | 2.4368
0.20144 | -33.9343 | 2.2973
0.20754 | -33.9315 | 2.1793
0.21365 | -33.9305 | 2.0669
0.21975 | -33.93 | 1.9707
0.22586 | -33.9309 | 1.8846
0.23196 | -33.9331 | 1.8129
0.23807 | -33.9368 | 1.7503
0.24417 | -33.9406 | 1.6916
0.25027 | -33.9443 | 1.6294
0.25638 | -33.9465 | 1.5677
0.26248 | -33.9482 | 1.5016
0.26859 | -33.949 | 1.4369
0.27469 | -33.9492 | 1.3667
0.2808 | -33.9501 | 1.2944
0.2869 | -33.9507 | 1.2246
0.293 | -33.9515 | 1.1585
0.29911 | -33.9523 | 1.0992
0.30521 | -33.9531 | 1.0409
0.31132 | -33.9537 | 1.002
0.31742 | -33.9547 | 0.95918
0.32353 | -33.955 | 0.9239
0.32963 | -33.9557 | 0.89429
0.33573 | -33.9566 | 0.87305
0.34184 | -33.9578 | 0.85081
0.34794 | -33.9597 | 0.8266
0.35405 | -33.9617 | 0.80205
0.36015 | -33.9639 | 0.77636
0.36626 | -33.9658 | 0.74865
0.37236 | -33.9677 | 0.71225
0.37846 | -33.9697 | 0.6687
0.38457 | -33.9706 | 0.61286
0.39067 | -33.9715 | 0.56696
0.39678 | -33.9721 | 0.51611
0.40288 | -33.9724 | 0.4632
0.40899 | -33.9723 | 0.41235
0.41509 | -33.9716 | 0.36251
0.42119 | -33.9715 | 0.32001
0.4273 | -33.9714 | 0.27467
0.4334 | -33.9712 | 0.22949
0.43951 | -33.9714 | 0.18699
0.44561 | -33.9722 | 0.14801
0.45172 | -33.9736 | 0.11678
0.45782 | -33.9743 | 0.080373
0.46392 | -33.9754 | 0.044813
0.47003 | -33.9764 | 0.012898
0.47613 | -33.9771 | -0.020652
0.48224 | -33.9776 | -0.054871
0.48834 | -33.9767 | -0.089092
0.49445 | -33.9765 | -0.12376
0.50055 | -33.976 | -0.15543
0.50665 | -33.9752 | -0.18912
0.51276 | -33.9743 | -0.2227
0.51886 | -33.9733 | -0.25616
0.52497 | -33.9721 | -0.28027
0.53107 | -33.9713 | -0.30943
0.53717 | -33.9706 | -0.33737
0.54328 | -33.9698 | -0.36017
0.54938 | -33.9695 | -0.38261
0.55549 | -33.9695 | -0.40353
0.56159 | -33.9705 | -0.41553
0.5677 | -33.9707 | -0.43041
0.5738 | -33.9717 | -0.4469
0.5799 | -33.9723 | -0.4621
0.58601 | -33.9733 | -0.47272
0.59211 | -33.9749 | -0.47969
0.59822 | -33.9761 | -0.48972
0.60432 | -33.9775 | -0.5022
0.61043 | -33.9787 | -0.51077
0.61653 | -33.9798 | -0.52153
0.62263 | -33.9812 | -0.53144
0.62874 | -33.9819 | -0.5426
0.63484 | -33.9826 | -0.55243
0.64095 | -33.9831 | -0.56315
0.64705 | -33.9834 | -0.57343
0.65316 | -33.9836 | -0.58626
0.65926 | -33.9837 | -0.60221
0.66536 | -33.9826 | -0.61963
0.67147 | -33.9822 | -0.63311
0.67757 | -33.9818 | -0.64758
0.68368 | -33.9811 | -0.6664
0.68978 | -33.9803 | -0.68362
0.69589 | -33.9794 | -0.70248
0.70199 | -33.9784 | -0.72485
0.70809 | -33.9776 | -0.74389
0.7142 | -33.9769 | -0.7641
0.7203 | -33.9758 | -0.78313
0.72641 | -33.9751 | -0.80525
0.73251 | -33.9745 | -0.825
0.73862 | -33.9747 | -0.84639
0.74472 | -33.9743 | -0.86615
0.75082 | -33.974 | -0.88789
0.75693 | -33.9743 | -0.90935
0.76303 | -33.9748 | -0.93098
0.76914 | -33.9765 | -0.95114
0.77524 | -33.9774 | -0.97051
0.78135 | -33.9786 | -0.99029
0.78745 | -33.9799 | -1.01
0.79355 | -33.981 | -1.0295
0.79966 | -33.9826 | -1.0486
0.80576 | -33.9848 | -1.0681
0.81187 | -33.9863 | -1.0892
0.81797 | -33.9881 | -1.1101
0.82408 | -33.9897 | -1.1303
0.83018 | -33.9914 | -1.1505
0.83628 | -33.9931 | -1.1688
0.84239 | -33.9943 | -1.189
0.84849 | -33.9956 | -1.2102
0.8546 | -33.997 | -1.2298
0.8607 | -33.9981 | -1.2501
0.86681 | -33.9992 | -1.2688
0.87291 | -34.0001 | -1.2874
0.87901 | -34.0005 | -1.3052
0.88512 | -34.0012 | -1.3231
0.89122 | -34.0019 | -1.3412
0.89733 | -34.0022 | -1.3567
0.90343 | -34.0026 | -1.3741
0.90953 | -34.0028 | -1.3885
0.91564 | -34.0028 | -1.4042
0.92174 | -34.003 | -1.4191
0.92785 | -34.003 | -1.4337
0.93395 | -34.0031 | -1.4475
0.94006 | -34.0033 | -1.4604
0.94616 | -34.003 | -1.4724
0.95226 | -34.0031 | -1.4851
0.95837 | -34.0031 | -1.499
0.96447 | -34.0033 | -1.5118
0.97058 | -34.0031 | -1.5252
0.97668 | -34.0029 | -1.5386
0.98279 | -34.0026 | -1.5513
0.98889 | -34.0024 | -1.5621
0.99499 | -34.0024 | -1.5755
1.0011 | -34.0022 | -1.5863
1.0072 | -34.0019 | -1.598
1.0133 | -34.0015 | -1.6088
1.0194 | -34.0011 | -1.6217
1.0255 | -34.0008 | -1.6354
1.0316 | -34.0003 | -1.6473
1.0377 | -33.9999 | -1.6609
1.0438 | -33.9993 | -1.6736
1.0499 | -33.9986 | -1.6887
1.056 | -33.9979 | -1.7029
1.0621 | -33.9973 | -1.7171
1.0682 | -33.9968 | -1.7295
1.0743 | -33.9962 | -1.7447
1.0805 | -33.9957 | -1.759
1.0866 | -33.9948 | -1.7739
1.0927 | -33.994 | -1.7882
1.0988 | -33.9934 | -1.8029
1.1049 | -33.9926 | -1.817
1.111 | -33.9918 | -1.8322
1.1171 | -33.9911 | -1.8466
1.1232 | -33.9907 | -1.8616
1.1293 | -33.9903 | -1.8754
1.1354 | -33.99 | -1.8902
1.1415 | -33.9896 | -1.9056
1.1476 | -33.989 | -1.9209
1.1537 | -33.9885 | -1.9354
1.1598 | -33.9886 | -1.95
1.1659 | -33.9883 | -1.9631
1.172 | -33.9881 | -1.9769
1.1781 | -33.9878 | -1.9907
1.1842 | -33.9877 | -2.0048
1.1903 | -33.988 | -2.0158
1.1964 | -33.9879 | -2.0291
1.2025 | -33.988 | -2.0429
1.2086 | -33.9883 | -2.0552
1.2147 | -33.9883 | -2.0676
1.2209 | -33.9886 | -2.0788
1.227 | -33.9893 | -2.0896
1.2331 | -33.9897 | -2.1015
1.2392 | -33.99 | -2.1133
1.2453 | -33.9904 | -2.1244
1.2514 | -33.9908 | -2.1344
1.2575 | -33.9914 | -2.1454
1.2636 | -33.9921 | -2.1545
1.2697 | -33.9927 | -2.1638
1.2758 | -33.9933 | -2.175
1.2819 | -33.994 | -2.185
1.288 | -33.9946 | -2.1957
1.2941 | -33.9954 | -2.2043
1.3002 | -33.996 | -2.2126
1.3063 | -33.9965 | -2.2217
1.3124 | -33.9972 | -2.2308
1.3185 | -33.9979 | -2.24
1.3246 | -33.9985 | -2.2492
1.3307 | -33.9993 | -2.2579
1.3368 | -34 | -2.2667
1.3429 | -34.0006 | -2.2758
1.349 | -34.0011 | -2.2852
1.3551 | -34.0018 | -2.2947
1.3613 | -34.0023 | -2.3031
1.3674 | -34.0029 | -2.3131
1.3735 | -34.0033 | -2.3229
1.3796 | -34.0037 | -2.332
1.3857 | -34.004 | -2.3413
1.3918 | -34.0045 | -2.3515
1.3979 | -34.005 | -2.3607
1.404 | -34.0055 | -2.3719
1.4101 | -34.0059 | -2.3822
1.4162 | -34.0062 | -2.3928
1.4223 | -34.0067 | -2.404
1.4284 | -34.0071 | -2.414
1.4345 | -34.0074 | -2.4255
1.4406 | -34.0076 | -2.4387
1.4467 | -34.0081 | -2.4502
1.4528 | -34.0085 | -2.4614
1.4589 | -34.0089 | -2.4731
1.465 | -34.0091 | -2.4838
1.4711 | -34.0096 | -2.495
1.4772 | -34.0101 | -2.5081
1.4833 | -34.0105 | -2.5209
1.4894 | -34.011 | -2.5329
1.4955 | -34.0115 | -2.5447
1.5016 | -34.012 | -2.557
1.5078 | -34.0123 | -2.5705
1.5139 | -34.0127 | -2.5834
1.52 | -34.0131 | -2.5957
1.5261 | -34.0134 | -2.6081
1.5322 | -34.0137 | -2.6218
1.5383 | -34.0141 | -2.6341
1.5444 | -34.0146 | -2.6465
1.5505 | -34.015 | -2.6597
1.5566 | -34.0155 | -2.6724
1.5627 | -34.0159 | -2.6859
1.5688 | -34.0165 | -2.6992
1.5749 | -34.0168 | -2.7115
1.581 | -34.0171 | -2.7246
1.5871 | -34.0174 | -2.7381
1.5932 | -34.0178 | -2.7507
1.5993 | -34.0181 | -2.7646
1.6054 | -34.0185 | -2.7775
1.6115 | -34.0189 | -2.7913
1.6176 | -34.0191 | -2.8037
1.6237 | -34.0193 | -2.817
1.6298 | -34.0195 | -2.8293
1.6359 | -34.0199 | -2.8418
1.642 | -34.0203 | -2.8535
1.6482 | -34.0204 | -2.8663
1.6543 | -34.0207 | -2.8797
1.6604 | -34.021 | -2.8915
1.6665 | -34.0211 | -2.9037
1.6726 | -34.0213 | -2.916
1.6787 | -34.0216 | -2.9283
1.6848 | -34.0215 | -2.9406
1.6909 | -34.0217 | -2.9531
1.697 | -34.0219 | -2.9651
1.7031 | -34.0219 | -2.9775
1.7092 | -34.0219 | -2.9894
1.7153 | -34.022 | -3.0012
1.7214 | -34.0218 | -3.0133
1.7275 | -34.0218 | -3.0256
1.7336 | -34.0217 | -3.0384
1.7397 | -34.0218 | -3.0505
1.7458 | -34.0218 | -3.0625
1.7519 | -34.0218 | -3.0745
1.758 | -34.0218 | -3.0862
1.7641 | -34.0219 | -3.0977
1.7702 | -34.0218 | -3.1104
1.7763 | -34.0218 | -3.1217
1.7824 | -34.0217 | -3.134
1.7885 | -34.0217 | -3.1463
1.7947 | -34.0218 | -3.1581
1.8008 | -34.0218 | -3.1693
1.8069 | -34.0218 | -3.1807
1.813 | -34.0218 | -3.1927
1.8191 | -34.0218 | -3.2053
1.8252 | -34.0218 | -3.2178
1.8313 | -34.0219 | -3.2299
1.8374 | -34.0221 | -3.2416
1.8435 | -34.0221 | -3.2534
1.8496 | -34.0221 | -3.2663
1.8557 | -34.0223 | -3.2789
1.8618 | -34.0225 | -3.291
1.8679 | -34.0226 | -3.3029
1.874 | -34.0227 | -3.3148
1.8801 | -34.0228 | -3.3271
1.8862 | -34.023 | -3.3389
1.8923 | -34.0231 | -3.351
1.8984 | -34.0232 | -3.363
1.9045 | -34.0233 | -3.3745
1.9106 | -34.0233 | -3.3866
1.9167 | -34.0233 | -3.3987
1.9228 | -34.0234 | -3.4104
1.9289 | -34.0235 | -3.4223
1.9351 | -34.0237 | -3.4331
1.9412 | -34.0239 | -3.4443
1.9473 | -34.0241 | -3.4563
1.9534 | -34.0243 | -3.4677
1.9595 | -34.0244 | -3.4794
1.9656 | -34.0247 | -3.4917
1.9717 | -34.0249 | -3.5033
1.9778 | -34.0252 | -3.5146
1.9839 | -34.0254 | -3.5261
1.99 | -34.0255 | -3.5372
1.9961 | -34.0256 | -3.5492
2.0022 | -34.0258 | -3.5606
2.0083 | -34.0259 | -3.5718
2.0144 | -34.0261 | -3.5832
2.0205 | -34.0264 | -3.5949
2.0266 | -34.0267 | -3.6062
2.0327 | -34.0269 | -3.6178
2.0388 | -34.0273 | -3.6293
2.0449 | -34.0274 | -3.6407
2.051 | -34.0276 | -3.652
2.0571 | -34.0279 | -3.6633
2.0632 | -34.0281 | -3.6747
2.0693 | -34.0283 | -3.6868
2.0754 | -34.0285 | -3.6981
2.0816 | -34.0288 | -3.7101
2.0877 | -34.029 | -3.7216
2.0938 | -34.0293 | -3.7336
2.0999 | -34.0295 | -3.7453
2.106 | -34.0297 | -3.7568
2.1121 | -34.0298 | -3.7687
2.1182 | -34.0301 | -3.7794
2.1243 | -34.0303 | -3.7902
2.1304 | -34.0305 | -3.8013
2.1365 | -34.0308 | -3.813
2.1426 | -34.031 | -3.8248
2.1487 | -34.0313 | -3.836
2.1548 | -34.0317 | -3.8477
2.1609 | -34.0321 | -3.8591
2.167 | -34.0324 | -3.8703
2.1731 | -34.0327 | -3.8822
2.1792 | -34.0329 | -3.8932
2.1853 | -34.0333 | -3.905
2.1914 | -34.0336 | -3.9163
2.1975 | -34.0339 | -3.9278
2.2036 | -34.0341 | -3.9398
2.2097 | -34.0344 | -3.9512
2.2158 | -34.0347 | -3.9618
2.222 | -34.035 | -3.9723
2.2281 | -34.0352 | -3.9836
2.2342 | -34.0355 | -3.9949
2.2403 | -34.0357 | -4.0055
2.2464 | -34.036 | -4.0162
2.2525 | -34.0362 | -4.0275
2.2586 | -34.0365 | -4.0381
2.2647 | -34.0368 | -4.0485
2.2708 | -34.0371 | -4.0597
2.2769 | -34.0374 | -4.0703
2.283 | -34.0376 | -4.0814
2.2891 | -34.0378 | -4.0923
2.2952 | -34.038 | -4.1028
2.3013 | -34.0382 | -4.1136
2.3074 | -34.0385 | -4.1241
2.3135 | -34.0387 | -4.1354
2.3196 | -34.0389 | -4.1456
2.3257 | -34.0391 | -4.1562
2.3318 | -34.0392 | -4.1673
2.3379 | -34.0394 | -4.178
2.344 | -34.0395 | -4.1894
2.3501 | -34.0397 | -4.2001
2.3562 | -34.0399 | -4.2104
2.3623 | -34.04 | -4.2208
2.3685 | -34.0401 | -4.2313
2.3746 | -34.0402 | -4.2425
2.3807 | -34.0405 | -4.2529
2.3868 | -34.0406 | -4.2636
2.3929 | -34.0408 | -4.2743
2.399 | -34.0409 | -4.2847
2.4051 | -34.0411 | -4.2952
2.4112 | -34.0413 | -4.3058
2.4173 | -34.0415 | -4.3165
2.4234 | -34.0415 | -4.3272
2.4295 | -34.0417 | -4.338
2.4356 | -34.0419 | -4.3488
2.4417 | -34.042 | -4.3599
2.4478 | -34.0421 | -4.3707
2.4539 | -34.0423 | -4.381
2.46 | -34.0423 | -4.3923
2.4661 | -34.0425 | -4.4033
2.4722 | -34.0426 | -4.4135
2.4783 | -34.0427 | -4.4239
2.4844 | -34.0428 | -4.4349
2.4905 | -34.0429 | -4.4458
2.4966 | -34.043 | -4.4565
2.5027 | -34.0431 | -4.467
2.5089 | -34.0432 | -4.4778
2.515 | -34.0433 | -4.4886
2.5211 | -34.0433 | -4.4993
2.5272 | -34.0434 | -4.5101
2.5333 | -34.0435 | -4.5211
2.5394 | -34.0437 | -4.5318
2.5455 | -34.0438 | -4.5431
2.5516 | -34.0439 | -4.5534
2.5577 | -34.0441 | -4.5647
2.5638 | -34.0442 | -4.5763
2.5699 | -34.0442 | -4.5877
2.576 | -34.0443 | -4.5991
2.5821 | -34.0444 | -4.6103
2.5882 | -34.0445 | -4.6213
2.5943 | -34.0446 | -4.6321
2.6004 | -34.0448 | -4.643
2.6065 | -34.0449 | -4.6533
2.6126 | -34.045 | -4.6637
2.6187 | -34.0451 | -4.6741
2.6248 | -34.0452 | -4.6848
2.6309 | -34.0453 | -4.6962
2.637 | -34.0456 | -4.7076
2.6431 | -34.0457 | -4.718
2.6492 | -34.0458 | -4.7293
2.6554 | -34.046 | -4.7401
2.6615 | -34.046 | -4.7509
2.6676 | -34.0463 | -4.7622
2.6737 | -34.0465 | -4.7733
2.6798 | -34.0468 | -4.7843
2.6859 | -34.0469 | -4.7954
2.692 | -34.0471 | -4.8062
2.6981 | -34.0472 | -4.8172
2.7042 | -34.0474 | -4.8283
2.7103 | -34.0476 | -4.8388
2.7164 | -34.0479 | -4.8497
2.7225 | -34.0481 | -4.8609
2.7286 | -34.0483 | -4.8718
2.7347 | -34.0485 | -4.8823
2.7408 | -34.0487 | -4.8923
2.7469 | -34.0489 | -4.9035
2.753 | -34.0492 | -4.914
2.7591 | -34.0494 | -4.9252
2.7652 | -34.0496 | -4.9365
2.7713 | -34.0498 | -4.9472
2.7774 | -34.05 | -4.958
2.7835 | -34.0502 | -4.9689
2.7896 | -34.0504 | -4.9793
2.7958 | -34.0506 | -4.99
2.8019 | -34.0509 | -5.0003
2.808 | -34.0511 | -5.011
2.8141 | -34.0514 | -5.022
2.8202 | -34.0517 | -5.0328
2.8263 | -34.052 | -5.0436
2.8324 | -34.0523 | -5.0548
2.8385 | -34.0526 | -5.0655
2.8446 | -34.053 | -5.0762
2.8507 | -34.0533 | -5.0869
2.8568 | -34.0536 | -5.0973
2.8629 | -34.0539 | -5.1078
2.869 | -34.0542 | -5.1188
2.8751 | -34.0545 | -5.1295
2.8812 | -34.0549 | -5.1406
2.8873 | -34.0551 | -5.1514
2.8934 | -34.0554 | -5.1623
2.8995 | -34.0557 | -5.1733
2.9056 | -34.056 | -5.1844
2.9117 | -34.0563 | -5.1953
2.9178 | -34.0566 | -5.2059
2.9239 | -34.057 | -5.2173
2.93 | -34.0574 | -5.2286
2.9361 | -34.0577 | -5.2393
2.9423 | -34.058 | -5.2499
2.9484 | -34.0584 | -5.2609
2.9545 | -34.0587 | -5.2718
2.9606 | -34.0591 | -5.283
2.9667 | -34.0594 | -5.2941
2.9728 | -34.0597 | -5.3053
2.9789 | -34.0602 | -5.3162
2.985 | -34.0606 | -5.3274
2.9911 | -34.0609 | -5.3386
2.9972 | -34.0613 | -5.3496
3.0033 | -34.0617 | -5.361
3.0094 | -34.0621 | -5.3719
3.0155 | -34.0625 | -5.3828
3.0216 | -34.0628 | -5.3939
3.0277 | -34.0631 | -5.4055
3.0338 | -34.0635 | -5.4167
3.0399 | -34.0639 | -5.4279
3.046 | -34.0642 | -5.4394
3.0521 | -34.0645 | -5.4504
3.0582 | -34.0648 | -5.4623
3.0643 | -34.0651 | -5.4737
3.0704 | -34.0654 | -5.4854
3.0765 | -34.0657 | -5.4964
3.0827 | -34.066 | -5.5075
3.0888 | -34.0663 | -5.5187
3.0949 | -34.0667 | -5.5299
3.101 | -34.067 | -5.5413
3.1071 | -34.0673 | -5.5529
3.1132 | -34.0677 | -5.5646
3.1193 | -34.068 | -5.576
3.1254 | -34.0684 | -5.5872
3.1315 | -34.0687 | -5.599
3.1376 | -34.069 | -5.6103
3.1437 | -34.0693 | -5.6217
3.1498 | -34.0697 | -5.6337
3.1559 | -34.07 | -5.645
3.162 | -34.0702 | -5.657
3.1681 | -34.0705 | -5.6688
3.1742 | -34.0708 | -5.6803
3.1803 | -34.0711 | -5.6918
3.1864 | -34.0714 | -5.703
3.1925 | -34.0717 | -5.7145
3.1986 | -34.072 | -5.7263
3.2047 | -34.0722 | -5.7383
3.2108 | -34.0725 | -5.7493
3.2169 | -34.0728 | -5.7609
3.223 | -34.073 | -5.7726
3.2292 | -34.0733 | -5.7842
3.2353 | -34.0735 | -5.795
3.2414 | -34.0737 | -5.8064
3.2475 | -34.0739 | -5.8177
3.2536 | -34.0742 | -5.8291
3.2597 | -34.0744 | -5.84
3.2658 | -34.0746 | -5.8508
3.2719 | -34.0748 | -5.8618
3.278 | -34.0751 | -5.8733
3.2841 | -34.0753 | -5.8843
3.2902 | -34.0756 | -5.8957
3.2963 | -34.0758 | -5.9071
3.3024 | -34.0759 | -5.9181
3.3085 | -34.0762 | -5.9296
3.3146 | -34.0763 | -5.9407
3.3207 | -34.0766 | -5.952
3.3268 | -34.0768 | -5.9634
3.3329 | -34.077 | -5.9742
3.339 | -34.0771 | -5.9849
3.3451 | -34.0773 | -5.9961
3.3512 | -34.0775 | -6.0067
3.3573 | -34.0777 | -6.0179
3.3634 | -34.0779 | -6.029
3.3696 | -34.0781 | -6.0403
3.3757 | -34.0782 | -6.051
3.3818 | -34.0784 | -6.0618
3.3879 | -34.0786 | -6.0728
3.394 | -34.0787 | -6.0841
3.4001 | -34.0789 | -6.0949
3.4062 | -34.079 | -6.1058
3.4123 | -34.0792 | -6.1163
3.4184 | -34.0794 | -6.1275
3.4245 | -34.0796 | -6.1383
3.4306 | -34.0796 | -6.1489
3.4367 | -34.0798 | -6.1594
3.4428 | -34.08 | -6.1698
3.4489 | -34.0801 | -6.1807
3.455 | -34.0802 | -6.1915
3.4611 | -34.0804 | -6.2026
3.4672 | -34.0806 | -6.2136
3.4733 | -34.0807 | -6.2246
3.4794 | -34.0809 | -6.2354
3.4855 | -34.081 | -6.2457
3.4916 | -34.0812 | -6.2561
3.4977 | -34.0813 | -6.2673
3.5038 | -34.0815 | -6.2781
3.5099 | -34.0816 | -6.2884
3.5161 | -34.0817 | -6.2987
3.5222 | -34.0819 | -6.309
3.5283 | -34.0819 | -6.3198
3.5344 | -34.0821 | -6.3305
3.5405 | -34.0822 | -6.341
3.5466 | -34.0824 | -6.3519
3.5527 | -34.0825 | -6.3626
3.5588 | -34.0826 | -6.3731
3.5649 | -34.0826 | -6.3837
3.571 | -34.0828 | -6.3943
3.5771 | -34.0829 | -6.4054
3.5832 | -34.083 | -6.4152
3.5893 | -34.0832 | -6.4258
3.5954 | -34.0834 | -6.4366
3.6015 | -34.0836 | -6.4474
3.6076 | -34.0838 | -6.4579
3.6137 | -34.0839 | -6.468
3.6198 | -34.0841 | -6.4782
3.6259 | -34.0843 | -6.489
3.632 | -34.0844 | -6.4996
3.6381 | -34.0846 | -6.5103
3.6442 | -34.0847 | -6.5208
3.6503 | -34.0849 | -6.5309
3.6565 | -34.0851 | -6.5416
3.6626 | -34.0852 | -6.5522
3.6687 | -34.0854 | -6.5628
3.6748 | -34.0857 | -6.5732
3.6809 | -34.0859 | -6.5838
3.687 | -34.0861 | -6.5947
3.6931 | -34.0863 | -6.6049
3.6992 | -34.0865 | -6.6154
3.7053 | -34.0867 | -6.6259
3.7114 | -34.0868 | -6.6365
3.7175 | -34.0871 | -6.6472
3.7236 | -34.0873 | -6.6572
3.7297 | -34.0875 | -6.6679
3.7358 | -34.0877 | -6.6787
3.7419 | -34.0879 | -6.6895
3.748 | -34.0881 | -6.7
3.7541 | -34.0883 | -6.7106
3.7602 | -34.0885 | -6.721
3.7663 | -34.0887 | -6.7314
3.7724 | -34.089 | -6.742
3.7785 | -34.0892 | -6.7524
3.7846 | -34.0894 | -6.7628
3.7907 | -34.0897 | -6.7736
3.7969 | -34.09 | -6.7843
3.803 | -34.0902 | -6.7944
3.8091 | -34.0904 | -6.8049
3.8152 | -34.0907 | -6.8153
3.8213 | -34.0909 | -6.8261
3.8274 | -34.0911 | -6.8364
3.8335 | -34.0915 | -6.847
3.8396 | -34.0917 | -6.8576
3.8457 | -34.092 | -6.8682
3.8518 | -34.0924 | -6.8787
3.8579 | -34.0927 | -6.8892
3.864 | -34.0929 | -6.8999
3.8701 | -34.0932 | -6.9108
3.8762 | -34.0935 | -6.9212
3.8823 | -34.0937 | -6.9319
3.8884 | -34.094 | -6.9425
3.8945 | -34.0944 | -6.9534
3.9006 | -34.0947 | -6.9643
3.9067 | -34.0949 | -6.9752
3.9128 | -34.0952 | -6.9859
3.9189 | -34.0955 | -6.9963
3.925 | -34.0957 | -7.007
3.9311 | -34.096 | -7.0172
3.9372 | -34.0963 | -7.028
3.9434 | -34.0966 | -7.0386
3.9495 | -34.0969 | -7.0493
3.9556 | -34.0972 | -7.0598
3.9617 | -34.0975 | -7.07
3.9678 | -34.0978 | -7.0804
3.9739 | -34.0981 | -7.0909
3.98 | -34.0984 | -7.1017
3.9861 | -34.0986 | -7.1122
3.9922 | -34.0989 | -7.1228
3.9983 | -34.0992 | -7.133
4.0044 | -34.0995 | -7.1432
4.0105 | -34.0997 | -7.1538
4.0166 | -34.1001 | -7.1645
4.0227 | -34.1004 | -7.175
4.0288 | -34.1006 | -7.1861
4.0349 | -34.1008 | -7.1964
4.041 | -34.1011 | -7.2069
4.0471 | -34.1013 | -7.2176
4.0532 | -34.1016 | -7.2279
4.0593 | -34.1019 | -7.2386
4.0654 | -34.1022 | -7.2489
4.0715 | -34.1025 | -7.26
4.0776 | -34.1028 | -7.2705
4.0838 | -34.1031 | -7.2814
4.0899 | -34.1034 | -7.2917
4.096 | -34.1036 | -7.3022
4.1021 | -34.1039 | -7.3126
4.1082 | -34.1041 | -7.3229
4.1143 | -34.1044 | -7.3331
4.1204 | -34.1047 | -7.3434
4.1265 | -34.105 | -7.3536
4.1326 | -34.1052 | -7.3638
4.1387 | -34.1056 | -7.3742
4.1448 | -34.1058 | -7.3848
4.1509 | -34.106 | -7.3953
4.157 | -34.1063 | -7.4062
4.1631 | -34.1066 | -7.4169
4.1692 | -34.1068 | -7.4272
4.1753 | -34.1071 | -7.4376
4.1814 | -34.1073 | -7.4478
4.1875 | -34.1076 | -7.4585
4.1936 | -34.1079 | -7.4691
4.1997 | -34.1081 | -7.4792
4.2058 | -34.1083 | -7.4899
4.2119 | -34.1086 | -7.5002
4.218 | -34.1088 | -7.5104
4.2241 | -34.109 | -7.5207
4.2303 | -34.1093 | -7.531
4.2364 | -34.1095 | -7.5414
4.2425 | -34.1097 | -7.5518
4.2486 | -34.11 | -7.5624
4.2547 | -34.1101 | -7.5726
4.2608 | -34.1103 | -7.5831
4.2669 | -34.1105 | -7.5932
4.273 | -34.1108 | -7.6037
4.2791 | -34.111 | -7.6141
4.2852 | -34.1112 | -7.6247
4.2913 | -34.1114 | -7.6348
4.2974 | -34.1116 | -7.6454
4.3035 | -34.1119 | -7.6561
4.3096 | -34.112 | -7.6667
4.3157 | -34.1123 | -7.6768
4.3218 | -34.1125 | -7.6871
4.3279 | -34.1127 | -7.6975
4.334 | -34.1129 | -7.7078
4.3401 | -34.1131 | -7.7178
4.3462 | -34.1134 | -7.7282
4.3523 | -34.1136 | -7.7382
4.3584 | -34.1138 | -7.7481
4.3645 | -34.114 | -7.7584
4.3707 | -34.1142 | -7.7684
4.3768 | -34.1144 | -7.7787
4.3829 | -34.1146 | -7.7893
4.389 | -34.1148 | -7.7996
4.3951 | -34.115 | -7.8097
4.4012 | -34.1152 | -7.8203
4.4073 | -34.1154 | -7.8304
4.4134 | -34.1156 | -7.8406
4.4195 | -34.1159 | -7.8509
4.4256 | -34.1161 | -7.8609
4.4317 | -34.1163 | -7.8711
4.4378 | -34.1165 | -7.8811
4.4439 | -34.1167 | -7.8911
4.45 | -34.1168 | -7.9013
4.4561 | -34.117 | -7.9115
4.4622 | -34.1172 | -7.9218
4.4683 | -34.1174 | -7.932
4.4744 | -34.1176 | -7.9425
4.4805 | -34.1178 | -7.9524
4.4866 | -34.118 | -7.9626
4.4927 | -34.1182 | -7.973
4.4988 | -34.1184 | -7.9832
4.5049 | -34.1186 | -7.9931
4.511 | -34.1189 | -8.003
4.5172 | -34.1191 | -8.0131
4.5233 | -34.1192 | -8.0237
4.5294 | -34.1195 | -8.0337
4.5355 | -34.1196 | -8.044
4.5416 | -34.1198 | -8.0541
4.5477 | -34.12 | -8.0643
4.5538 | -34.1202 | -8.0741
4.5599 | -34.1205 | -8.0842
4.566 | -34.1207 | -8.0943
4.5721 | -34.1209 | -8.1045
4.5782 | -34.1211 | -8.1149
4.5843 | -34.1213 | -8.1252
4.5904 | -34.1216 | -8.1354
4.5965 | -34.1218 | -8.1455
4.6026 | -34.122 | -8.1561
4.6087 | -34.1222 | -8.1661
4.6148 | -34.1223 | -8.1766
4.6209 | -34.1226 | -8.1867
4.627 | -34.1228 | -8.1971
4.6331 | -34.123 | -8.2072
4.6392 | -34.1232 | -8.2175
4.6453 | -34.1234 | -8.2276
4.6514 | -34.1236 | -8.2378
4.6576 | -34.1238 | -8.2484
4.6637 | -34.1239 | -8.2582
4.6698 | -34.1241 | -8.2683
4.6759 | -34.1243 | -8.2788
4.682 | -34.1245 | -8.2888
4.6881 | -34.1247 | -8.299
4.6942 | -34.1249 | -8.3092
4.7003 | -34.1251 | -8.3194
4.7064 | -34.1253 | -8.3294
4.7125 | -34.1256 | -8.3398
4.7186 | -34.1257 | -8.3498
4.7247 | -34.126 | -8.36
4.7308 | -34.1262 | -8.3702
4.7369 | -34.1264 | -8.3805
4.743 | -34.1266 | -8.391
4.7491 | -34.1268 | -8.4015
4.7552 | -34.127 | -8.4118
4.7613 | -34.1272 | -8.4221
4.7674 | -34.1274 | -8.4322
4.7735 | -34.1276 | -8.4428
4.7796 | -34.1278 | -8.4533
4.7857 | -34.128 | -8.4636
4.7918 | -34.1283 | -8.4737
4.7979 | -34.1285 | -8.4838
4.8041 | -34.1287 | -8.4941
4.8102 | -34.1289 | -8.504
4.8163 | -34.1291 | -8.5141
4.8224 | -34.1294 | -8.5244
4.8285 | -34.1296 | -8.5344
4.8346 | -34.1298 | -8.5444
4.8407 | -34.13 | -8.5547
4.8468 | -34.1303 | -8.5652
4.8529 | -34.1305 | -8.5757
4.859 | -34.1307 | -8.5862
4.8651 | -34.1309 | -8.5962
4.8712 | -34.1312 | -8.6063
4.8773 | -34.1313 | -8.6166
4.8834 | -34.1316 | -8.6269
4.8895 | -34.1318 | -8.6369
4.8956 | -34.132 | -8.6471
4.9017 | -34.1322 | -8.6572
4.9078 | -34.1324 | -8.6677
4.9139 | -34.1326 | -8.6776
4.92 | -34.1329 | -8.6875
4.9261 | -34.1331 | -8.6974
4.9322 | -34.1334 | -8.7077
4.9383 | -34.1336 | -8.7179
4.9445 | -34.1339 | -8.7282
4.9506 | -34.1341 | -8.7382
4.9567 | -34.1343 | -8.7482
4.9628 | -34.1345 | -8.7584
4.9689 | -34.1347 | -8.769
4.975 | -34.1349 | -8.7795
4.9811 | -34.1352 | -8.7894
4.9872 | -34.1355 | -8.7996
4.9933 | -34.1358 | -8.81
4.9994 | -34.136 | -8.8206
5.0055 | -34.1363 | -8.8309
5.0116 | -34.1365 | -8.8416
5.0177 | -34.1368 | -8.8521
5.0238 | -34.1371 | -8.8626
5.0299 | -34.1374 | -8.8732
5.036 | -34.1376 | -8.8835
5.0421 | -34.1378 | -8.8936
5.0482 | -34.1381 | -8.9039
5.0543 | -34.1384 | -8.914
5.0604 | -34.1386 | -8.9242
5.0665 | -34.1389 | -8.9342
5.0726 | -34.1392 | -8.9448
5.0787 | -34.1394 | -8.9553
5.0848 | -34.1397 | -8.9656
5.091 | -34.1399 | -8.9757
5.0971 | -34.1402 | -8.9857
5.1032 | -34.1405 | -8.9963
5.1093 | -34.1408 | -9.0065
5.1154 | -34.141 | -9.0166
5.1215 | -34.1412 | -9.0266
5.1276 | -34.1415 | -9.0367
5.1337 | -34.1418 | -9.0467
5.1398 | -34.142 | -9.0567
5.1459 | -34.1423 | -9.0672
5.152 | -34.1426 | -9.0777
5.1581 | -34.1428 | -9.0882
5.1642 | -34.1431 | -9.0985
5.1703 | -34.1434 | -9.1088
5.1764 | -34.1437 | -9.1189
5.1825 | -34.1439 | -9.1289
5.1886 | -34.1442 | -9.1389
5.1947 | -34.1444 | -9.1492
5.2008 | -34.1447 | -9.1596
5.2069 | -34.145 | -9.1701
5.213 | -34.1453 | -9.1802
5.2191 | -34.1456 | -9.1905
5.2252 | -34.1458 | -9.2009
5.2314 | -34.1461 | -9.2113
5.2375 | -34.1463 | -9.2214
5.2436 | -34.1466 | -9.2317
5.2497 | -34.1469 | -9.2418
5.2558 | -34.1472 | -9.2524
5.2619 | -34.1474 | -9.2629
5.268 | -34.1477 | -9.2734
5.2741 | -34.1481 | -9.2837
5.2802 | -34.1483 | -9.2944
5.2863 | -34.1486 | -9.3046
5.2924 | -34.1489 | -9.3151
5.2985 | -34.1492 | -9.3255
5.3046 | -34.1495 | -9.3359
5.3107 | -34.1498 | -9.3464
5.3168 | -34.1501 | -9.3568
5.3229 | -34.1503 | -9.3671
5.329 | -34.1506 | -9.3775
5.3351 | -34.1509 | -9.3875
5.3412 | -34.1512 | -9.3978
5.3473 | -34.1515 | -9.4082
5.3534 | -34.1518 | -9.4187
5.3595 | -34.1521 | -9.4291
5.3656 | -34.1524 | -9.4394
5.3717 | -34.1526 | -9.4501
5.3779 | -34.1529 | -9.4608
5.384 | -34.1532 | -9.4711
5.3901 | -34.1535 | -9.4817
5.3962 | -34.1538 | -9.4919
5.4023 | -34.1541 | -9.5025
5.4084 | -34.1544 | -9.5127
5.4145 | -34.1548 | -9.5229
5.4206 | -34.155 | -9.5333
5.4267 | -34.1553 | -9.544
5.4328 | -34.1556 | -9.5544
5.4389 | -34.1559 | -9.5645
5.445 | -34.1562 | -9.5753
5.4511 | -34.1565 | -9.5854
5.4572 | -34.1568 | -9.5959
5.4633 | -34.1571 | -9.6062
5.4694 | -34.1573 | -9.6166
5.4755 | -34.1576 | -9.6268
5.4816 | -34.1579 | -9.6372
5.4877 | -34.1582 | -9.6474
5.4938 | -34.1585 | -9.6575
5.4999 | -34.1588 | -9.6678
5.506 | -34.1592 | -9.6782
5.5121 | -34.1594 | -9.6884
5.5183 | -34.1597 | -9.6987
5.5244 | -34.16 | -9.7093
5.5305 | -34.1603 | -9.7196
5.5366 | -34.1606 | -9.7297
5.5427 | -34.1609 | -9.7402
5.5488 | -34.1611 | -9.7506
5.5549 | -34.1614 | -9.7611
5.561 | -34.1617 | -9.7717
5.5671 | -34.1621 | -9.7819
5.5732 | -34.1623 | -9.7925
5.5793 | -34.1627 | -9.8032
5.5854 | -34.163 | -9.8135
5.5915 | -34.1633 | -9.8242
5.5976 | -34.1636 | -9.8347
5.6037 | -34.1638 | -9.8451
5.6098 | -34.1641 | -9.8554
5.6159 | -34.1644 | -9.8657
5.622 | -34.1647 | -9.8761
5.6281 | -34.165 | -9.8866
5.6342 | -34.1653 | -9.8973
5.6403 | -34.1656 | -9.9077
5.6464 | -34.1659 | -9.9183
5.6525 | -34.1662 | -9.9288
5.6586 | -34.1665 | -9.9389
5.6648 | -34.1669 | -9.9493
5.6709 | -34.1672 | -9.9599
5.677 | -34.1675 | -9.9703
5.6831 | -34.1678 | -9.9808
5.6892 | -34.1681 | -9.9914
5.6953 | -34.1684 | -10.0018
5.7014 | -34.1687 | -10.0123
5.7075 | -34.1689 | -10.0227
5.7136 | -34.1692 | -10.0331
5.7197 | -34.1695 | -10.0435
5.7258 | -34.1698 | -10.0537
5.7319 | -34.1701 | -10.0641
5.738 | -34.1704 | -10.0745
5.7441 | -34.1708 | -10.0849
5.7502 | -34.1711 | -10.095
5.7563 | -34.1714 | -10.1054
5.7624 | -34.1717 | -10.116
5.7685 | -34.172 | -10.1265
5.7746 | -34.1723 | -10.1372
5.7807 | -34.1725 | -10.1476
5.7868 | -34.1728 | -10.1582
5.7929 | -34.1731 | -10.1685
5.799 | -34.1734 | -10.1789
5.8052 | -34.1737 | -10.1891
5.8113 | -34.174 | -10.1994
5.8174 | -34.1743 | -10.2097
5.8235 | -34.1746 | -10.2201
5.8296 | -34.1748 | -10.2307
5.8357 | -34.1751 | -10.241
5.8418 | -34.1754 | -10.2515
5.8479 | -34.1757 | -10.262
5.854 | -34.176 | -10.2726
5.8601 | -34.1763 | -10.2828
5.8662 | -34.1766 | -10.2934
5.8723 | -34.1769 | -10.3036
5.8784 | -34.1771 | -10.3139
5.8845 | -34.1774 | -10.3242
5.8906 | -34.1777 | -10.3347
5.8967 | -34.178 | -10.3449
5.9028 | -34.1783 | -10.3551
5.9089 | -34.1787 | -10.3653
5.915 | -34.1789 | -10.3755
5.9211 | -34.1792 | -10.3858
5.9272 | -34.1795 | -10.396
5.9333 | -34.1798 | -10.4063
5.9394 | -34.1801 | -10.4166
5.9455 | -34.1804 | -10.4268
5.9517 | -34.1808 | -10.437
5.9578 | -34.1811 | -10.4474
5.9639 | -34.1814 | -10.4575
5.97 | -34.1818 | -10.4679
5.9761 | -34.1821 | -10.4782
5.9822 | -34.1823 | -10.4886
5.9883 | -34.1827 | -10.4989
5.9944 | -34.183 | -10.5093
6.0005 | -34.1833 | -10.5198
6.0066 | -34.1836 | -10.5299
6.0127 | -34.1839 | -10.54
6.0188 | -34.1842 | -10.5506
6.0249 | -34.1845 | -10.561
6.031 | -34.1848 | -10.5712
6.0371 | -34.1851 | -10.5815
6.0432 | -34.1854 | -10.5917
6.0493 | -34.1857 | -10.6018
6.0554 | -34.186 | -10.612
6.0615 | -34.1863 | -10.6223
6.0676 | -34.1866 | -10.6328
6.0737 | -34.187 | -10.6432
6.0798 | -34.1872 | -10.6535
6.0859 | -34.1876 | -10.6637
6.0921 | -34.1879 | -10.6738
6.0982 | -34.1882 | -10.6841
6.1043 | -34.1885 | -10.6943
6.1104 | -34.1888 | -10.7048
6.1165 | -34.1892 | -10.7146
6.1226 | -34.1895 | -10.7248
6.1287 | -34.1898 | -10.735
6.1348 | -34.1901 | -10.7453
6.1409 | -34.1904 | -10.7555
6.147 | -34.1907 | -10.7657
6.1531 | -34.191 | -10.7759
6.1592 | -34.1913 | -10.786
6.1653 | -34.1916 | -10.7962
6.1714 | -34.1919 | -10.8064
6.1775 | -34.1922 | -10.8168
6.1836 | -34.1925 | -10.8271
6.1897 | -34.1928 | -10.8373
6.1958 | -34.1931 | -10.8478
6.2019 | -34.1935 | -10.8581
6.208 | -34.1938 | -10.8684
6.2141 | -34.1941 | -10.8785
6.2202 | -34.1944 | -10.8889
6.2263 | -34.1948 | -10.899
6.2325 | -34.195 | -10.9089
6.2386 | -34.1954 | -10.9189
6.2447 | -34.1957 | -10.9293
6.2508 | -34.196 | -10.9395
6.2569 | -34.1964 | -10.9497
6.263 | -34.1966 | -10.9599
6.2691 | -34.1969 | -10.97
6.2752 | -34.1972 | -10.9803
6.2813 | -34.1975 | -10.9905
6.2874 | -34.1978 | -11.0006
6.2935 | -34.1981 | -11.0107
6.2996 | -34.1984 | -11.0209
6.3057 | -34.1987 | -11.0311
6.3118 | -34.199 | -11.0414
6.3179 | -34.1993 | -11.0516
6.324 | -34.1996 | -11.0617
6.3301 | -34.1999 | -11.0719
6.3362 | -34.2003 | -11.0819
6.3423 | -34.2006 | -11.0922
6.3484 | -34.2009 | -11.1025
6.3545 | -34.2012 | -11.1123
6.3606 | -34.2015 | -11.1223
6.3667 | -34.2018 | -11.1323
6.3728 | -34.2021 | -11.1425
6.379 | -34.2024 | -11.1527
6.3851 | -34.2027 | -11.1626
6.3912 | -34.203 | -11.1727
6.3973 | -34.2033 | -11.1828
6.4034 | -34.2036 | -11.1928
6.4095 | -34.2039 | -11.2029
6.4156 | -34.2042 | -11.2131
6.4217 | -34.2045 | -11.2232
6.4278 | -34.2048 | -11.2334
6.4339 | -34.2051 | -11.2434
6.44 | -34.2055 | -11.2535
6.4461 | -34.2057 | -11.2638
6.4522 | -34.2061 | -11.2737
6.4583 | -34.2063 | -11.284
6.4644 | -34.2066 | -11.2942
6.4705 | -34.2069 | -11.304
6.4766 | -34.2072 | -11.3141
6.4827 | -34.2075 | -11.3243
6.4888 | -34.2078 | -11.3342
6.4949 | -34.2081 | -11.3441
6.501 | -34.2084 | -11.3542
6.5071 | -34.2087 | -11.3641
6.5132 | -34.2091 | -11.3742
6.5194 | -34.2093 | -11.3843
6.5255 | -34.2097 | -11.3942
6.5316 | -34.2099 | -11.4044
6.5377 | -34.2102 | -11.4144
6.5438 | -34.2105 | -11.4243
6.5499 | -34.2108 | -11.4343
6.556 | -34.2111 | -11.4443
6.5621 | -34.2114 | -11.4546
6.5682 | -34.2117 | -11.4647
6.5743 | -34.2121 | -11.4747
6.5804 | -34.2124 | -11.4848
6.5865 | -34.2127 | -11.4949
6.5926 | -34.213 | -11.5049
6.5987 | -34.2133 | -11.5146
6.6048 | -34.2136 | -11.5247
6.6109 | -34.2139 | -11.5347
6.617 | -34.2142 | -11.5449
6.6231 | -34.2145 | -11.5549
6.6292 | -34.2148 | -11.5649
6.6353 | -34.2151 | -11.5752
6.6414 | -34.2155 | -11.5851
6.6475 | -34.2158 | -11.5951
6.6536 | -34.216 | -11.6051
6.6597 | -34.2163 | -11.6148
6.6659 | -34.2166 | -11.6249
6.672 | -34.217 | -11.6348
6.6781 | -34.2173 | -11.6448
6.6842 | -34.2176 | -11.6546
6.6903 | -34.2179 | -11.6646
6.6964 | -34.2183 | -11.6747
6.7025 | -34.2185 | -11.6845
6.7086 | -34.2188 | -11.6944
6.7147 | -34.2191 | -11.7045
6.7208 | -34.2194 | -11.7146
6.7269 | -34.2197 | -11.7246
6.733 | -34.22 | -11.7343
6.7391 | -34.2203 | -11.7443
6.7452 | -34.2207 | -11.7541
6.7513 | -34.221 | -11.7638
6.7574 | -34.2213 | -11.7738
6.7635 | -34.2216 | -11.7836
6.7696 | -34.2219 | -11.7934
6.7757 | -34.2222 | -11.8034
6.7818 | -34.2226 | -11.8134
6.7879 | -34.2229 | -11.8231
6.794 | -34.2232 | -11.8329
6.8001 | -34.2235 | -11.8428
6.8063 | -34.2238 | -11.8527
6.8124 | -34.2241 | -11.8628
6.8185 | -34.2244 | -11.8729
6.8246 | -34.2246 | -11.8826
6.8307 | -34.2249 | -11.8924
6.8368 | -34.2253 | -11.9025
6.8429 | -34.2256 | -11.9125
6.849 | -34.2259 | -11.9225
6.8551 | -34.2263 | -11.9322
6.8612 | -34.2266 | -11.9421
6.8673 | -34.2268 | -11.952
6.8734 | -34.2272 | -11.962
6.8795 | -34.2275 | -11.9719
6.8856 | -34.2278 | -11.9818
6.8917 | -34.2281 | -11.9916
6.8978 | -34.2284 | -12.0013
6.9039 | -34.2287 | -12.0109
6.91 | -34.229 | -12.0209
6.9161 | -34.2294 | -12.0305
6.9222 | -34.2296 | -12.0404
6.9283 | -34.23 | -12.05
6.9344 | -34.2303 | -12.0599
6.9405 | -34.2307 | -12.0697
6.9466 | -34.231 | -12.0798
6.9528 | -34.2313 | -12.0898
6.9589 | -34.2317 | -12.0998
6.965 | -34.232 | -12.1097
6.9711 | -34.2323 | -12.1197
6.9772 | -34.2327 | -12.1295
6.9833 | -34.233 | -12.1394
6.9894 | -34.2333 | -12.1495
6.9955 | -34.2336 | -12.1591
7.0016 | -34.2339 | -12.1689
7.0077 | -34.2343 | -12.1787
7.0138 | -34.2346 | -12.1885
7.0199 | -34.2349 | -12.1982
7.026 | -34.2352 | -12.2081
7.0321 | -34.2355 | -12.2177
7.0382 | -34.2359 | -12.2272
7.0443 | -34.2362 | -12.2372
7.0504 | -34.2365 | -12.2472
7.0565 | -34.2369 | -12.257
7.0626 | -34.2372 | -12.2666
7.0687 | -34.2375 | -12.276
7.0748 | -34.2379 | -12.2856
7.0809 | -34.2382 | -12.2954
7.087 | -34.2385 | -12.3052
7.0932 | -34.2388 | -12.3151
7.0993 | -34.2391 | -12.3248
7.1054 | -34.2395 | -12.3345
7.1115 | -34.2398 | -12.3441
7.1176 | -34.2401 | -12.3538
7.1237 | -34.2404 | -12.3635
7.1298 | -34.2408 | -12.3732
7.1359 | -34.2411 | -12.383
7.142 | -34.2414 | -12.3925
7.1481 | -34.2417 | -12.4022
7.1542 | -34.242 | -12.412
7.1603 | -34.2424 | -12.4217
7.1664 | -34.2427 | -12.4315
7.1725 | -34.2431 | -12.4413
7.1786 | -34.2434 | -12.451
7.1847 | -34.2437 | -12.4608
7.1908 | -34.244 | -12.4707
7.1969 | -34.2443 | -12.4807
7.203 | -34.2447 | -12.4902
7.2091 | -34.245 | -12.4999
7.2152 | -34.2454 | -12.5097
7.2213 | -34.2457 | -12.5193
7.2274 | -34.246 | -12.529
7.2335 | -34.2463 | -12.5387
7.2397 | -34.2467 | -12.5484
7.2458 | -34.247 | -12.5579
7.2519 | -34.2474 | -12.5676
7.258 | -34.2477 | -12.5775
7.2641 | -34.248 | -12.5873
7.2702 | -34.2483 | -12.5972
7.2763 | -34.2487 | -12.6069
7.2824 | -34.2491 | -12.6168
7.2885 | -34.2494 | -12.6266
7.2946 | -34.2497 | -12.6363
7.3007 | -34.25 | -12.6462
7.3068 | -34.2503 | -12.6562
7.3129 | -34.2507 | -12.6661
7.319 | -34.251 | -12.6758
7.3251 | -34.2513 | -12.6854
7.3312 | -34.2517 | -12.6951
7.3373 | -34.252 | -12.7046
7.3434 | -34.2524 | -12.7142
7.3495 | -34.2527 | -12.7239
7.3556 | -34.2531 | -12.7337
7.3617 | -34.2534 | -12.7436
7.3678 | -34.2537 | -12.7533
7.3739 | -34.2541 | -12.7632
7.3801 | -34.2544 | -12.7732
7.3862 | -34.2547 | -12.7829
7.3923 | -34.2551 | -12.7928
7.3984 | -34.2554 | -12.8023
7.4045 | -34.2558 | -12.812
7.4106 | -34.2562 | -12.8217
7.4167 | -34.2565 | -12.8315
7.4228 | -34.2569 | -12.8413
7.4289 | -34.2572 | -12.851
7.435 | -34.2575 | -12.8605
7.4411 | -34.2579 | -12.8704
7.4472 | -34.2582 | -12.8802
7.4533 | -34.2585 | -12.8897
7.4594 | -34.2589 | -12.8997
7.4655 | -34.2592 | -12.9093
7.4716 | -34.2596 | -12.9191
7.4777 | -34.2599 | -12.9288
7.4838 | -34.2603 | -12.9385
7.4899 | -34.2606 | -12.9482
7.496 | -34.2609 | -12.9578
7.5021 | -34.2613 | -12.9674
7.5082 | -34.2616 | -12.9772
7.5143 | -34.262 | -12.9869
7.5204 | -34.2623 | -12.9968
7.5266 | -34.2626 | -13.0064
7.5327 | -34.263 | -13.0161
7.5388 | -34.2633 | -13.0261
7.5449 | -34.2637 | -13.036
7.551 | -34.264 | -13.0456
7.5571 | -34.2644 | -13.0554
7.5632 | -34.2647 | -13.0653
7.5693 | -34.2651 | -13.0749
7.5754 | -34.2654 | -13.0846
7.5815 | -34.2658 | -13.0943
7.5876 | -34.2661 | -13.1043
7.5937 | -34.2665 | -13.1142
7.5998 | -34.2669 | -13.1239
7.6059 | -34.2672 | -13.1338
7.612 | -34.2676 | -13.1435
7.6181 | -34.2679 | -13.1532
7.6242 | -34.2683 | -13.1632
7.6303 | -34.2686 | -13.1732
7.6364 | -34.2689 | -13.183
7.6425 | -34.2693 | -13.1928
7.6486 | -34.2696 | -13.2026
7.6547 | -34.27 | -13.2122
7.6608 | -34.2703 | -13.2219
7.667 | -34.2707 | -13.2318
7.6731 | -34.271 | -13.2417
7.6792 | -34.2714 | -13.2517
7.6853 | -34.2717 | -13.2615
7.6914 | -34.2721 | -13.2714
7.6975 | -34.2724 | -13.2813
7.7036 | -34.2728 | -13.2911
7.7097 | -34.2732 | -13.3011
7.7158 | -34.2735 | -13.3109
7.7219 | -34.2739 | -13.3206
7.728 | -34.2742 | -13.3305
7.7341 | -34.2746 | -13.3404
7.7402 | -34.2749 | -13.3502
7.7463 | -34.2753 | -13.3601
7.7524 | -34.2756 | -13.3701
7.7585 | -34.276 | -13.38
7.7646 | -34.2763 | -13.3898
7.7707 | -34.2767 | -13.3998
7.7768 | -34.277 | -13.4098
7.7829 | -34.2774 | -13.4199
7.789 | -34.2777 | -13.4299
7.7951 | -34.278 | -13.4397
7.8012 | -34.2784 | -13.4496
7.8073 | -34.2788 | -13.4597
7.8135 | -34.2791 | -13.4695
7.8196 | -34.2795 | -13.4796
7.8257 | -34.2799 | -13.4896
7.8318 | -34.2802 | -13.4995
7.8379 | -34.2806 | -13.5093
7.844 | -34.2809 | -13.5191
7.8501 | -34.2813 | -13.5291
7.8562 | -34.2816 | -13.5392
7.8623 | -34.282 | -13.5489
7.8684 | -34.2823 | -13.5589
7.8745 | -34.2827 | -13.5688
7.8806 | -34.283 | -13.5789
7.8867 | -34.2834 | -13.5886
7.8928 | -34.2838 | -13.5982
7.8989 | -34.2841 | -13.6081
7.905 | -34.2845 | -13.6181
7.9111 | -34.2848 | -13.6281
7.9172 | -34.2852 | -13.6379
7.9233 | -34.2855 | -13.6479
7.9294 | -34.2859 | -13.6578
7.9355 | -34.2862 | -13.6675
7.9416 | -34.2865 | -13.6774
7.9477 | -34.2869 | -13.6872
7.9539 | -34.2872 | -13.697
7.96 | -34.2876 | -13.7071
7.9661 | -34.288 | -13.717
7.9722 | -34.2884 | -13.7267
7.9783 | -34.2887 | -13.7369
7.9844 | -34.2891 | -13.7468
7.9905 | -34.2895 | -13.7568
7.9966 | -34.2898 | -13.7669
8.0027 | -34.2902 | -13.7768
8.0088 | -34.2905 | -13.7866
8.0149 | -34.2908 | -13.7966
8.021 | -34.2912 | -13.8065
8.0271 | -34.2916 | -13.8167
8.0332 | -34.2919 | -13.8266
8.0393 | -34.2923 | -13.8366
8.0454 | -34.2926 | -13.8468
8.0515 | -34.293 | -13.8567
8.0576 | -34.2934 | -13.8668
8.0637 | -34.2937 | -13.8767
8.0698 | -34.2941 | -13.8865
8.0759 | -34.2944 | -13.8963
8.082 | -34.2947 | -13.9061
8.0881 | -34.2951 | -13.9161
8.0942 | -34.2955 | -13.926
8.1004 | -34.2959 | -13.936
8.1065 | -34.2962 | -13.9458
8.1126 | -34.2965 | -13.9557
8.1187 | -34.2969 | -13.9658
8.1248 | -34.2972 | -13.9755
8.1309 | -34.2976 | -13.9853
8.137 | -34.298 | -13.995
8.1431 | -34.2983 | -14.0052
8.1492 | -34.2987 | -14.015
8.1553 | -34.299 | -14.025
8.1614 | -34.2994 | -14.0351
8.1675 | -34.2997 | -14.045
8.1736 | -34.3001 | -14.0551
8.1797 | -34.3004 | -14.0651
8.1858 | -34.3008 | -14.0751
8.1919 | -34.3011 | -14.085
8.198 | -34.3015 | -14.0947
8.2041 | -34.3019 | -14.1047
8.2102 | -34.3022 | -14.1146
8.2163 | -34.3026 | -14.1244
8.2224 | -34.3029 | -14.1343
8.2285 | -34.3033 | -14.144
8.2346 | -34.3036 | -14.1539
8.2408 | -34.304 | -14.1636
8.2469 | -34.3044 | -14.1736
8.253 | -34.3048 | -14.1834
8.2591 | -34.3051 | -14.1934
8.2652 | -34.3055 | -14.2034
8.2713 | -34.3059 | -14.2133
8.2774 | -34.3062 | -14.2233
8.2835 | -34.3066 | -14.2335
8.2896 | -34.307 | -14.2433
8.2957 | -34.3073 | -14.2532
8.3018 | -34.3077 | -14.2632
8.3079 | -34.3081 | -14.2729
8.314 | -34.3084 | -14.2828
8.3201 | -34.3088 | -14.2926
8.3262 | -34.3092 | -14.3026
8.3323 | -34.3095 | -14.3125
8.3384 | -34.3099 | -14.3225
8.3445 | -34.3102 | -14.3323
8.3506 | -34.3106 | -14.3423
8.3567 | -34.311 | -14.3523
8.3628 | -34.3113 | -14.3622
8.3689 | -34.3117 | -14.3719
8.375 | -34.3121 | -14.3819
8.3812 | -34.3124 | -14.3917
8.3873 | -34.3128 | -14.4016
8.3934 | -34.3131 | -14.4116
8.3995 | -34.3135 | -14.4214
8.4056 | -34.3138 | -14.4312
8.4117 | -34.3142 | -14.4412
8.4178 | -34.3145 | -14.4512
8.4239 | -34.3149 | -14.4609
8.43 | -34.3152 | -14.4708
8.4361 | -34.3156 | -14.4808
8.4422 | -34.316 | -14.4906
8.4483 | -34.3163 | -14.5004
8.4544 | -34.3167 | -14.5103
8.4605 | -34.3171 | -14.52
8.4666 | -34.3174 | -14.5302
8.4727 | -34.3178 | -14.5401
8.4788 | -34.3181 | -14.5499
8.4849 | -34.3185 | -14.5598
8.491 | -34.3188 | -14.5697
8.4971 | -34.3192 | -14.5793
8.5032 | -34.3195 | -14.5891
8.5093 | -34.3199 | -14.5988
8.5154 | -34.3203 | -14.6087
8.5215 | -34.3207 | -14.6185
8.5277 | -34.321 | -14.6284
8.5338 | -34.3214 | -14.6381
8.5399 | -34.3217 | -14.648
8.546 | -34.3221 | -14.6577
8.5521 | -34.3225 | -14.6675
8.5582 | -34.3228 | -14.6773
8.5643 | -34.3232 | -14.6869
8.5704 | -34.3235 | -14.6965
8.5765 | -34.3239 | -14.7065
8.5826 | -34.3243 | -14.716
8.5887 | -34.3246 | -14.7258
8.5948 | -34.325 | -14.7356
8.6009 | -34.3254 | -14.7453
8.607 | -34.3257 | -14.7552
8.6131 | -34.3261 | -14.7651
8.6192 | -34.3265 | -14.7747
8.6253 | -34.3269 | -14.7845
8.6314 | -34.3272 | -14.7942
8.6375 | -34.3275 | -14.8041
8.6436 | -34.3279 | -14.8138
8.6497 | -34.3283 | -14.8235
8.6558 | -34.3286 | -14.8332
8.6619 | -34.3289 | -14.843
8.6681 | -34.3293 | -14.8529
8.6742 | -34.3297 | -14.8627
8.6803 | -34.33 | -14.8725
8.6864 | -34.3304 | -14.8819
8.6925 | -34.3308 | -14.8915
8.6986 | -34.3311 | -14.9014
8.7047 | -34.3315 | -14.9114
8.7108 | -34.3318 | -14.9211
8.7169 | -34.3322 | -14.9309
8.723 | -34.3326 | -14.9404
8.7291 | -34.3329 | -14.95
8.7352 | -34.3333 | -14.96
8.7413 | -34.3337 | -14.9698
8.7474 | -34.334 | -14.9793
8.7535 | -34.3344 | -14.9891
8.7596 | -34.3348 | -14.9988
8.7657 | -34.3351 | -15.0086
8.7718 | -34.3355 | -15.0184
8.7779 | -34.3359 | -15.0282
8.784 | -34.3362 | -15.038
8.7901 | -34.3366 | -15.0477
8.7962 | -34.337 | -15.0572
8.8023 | -34.3373 | -15.0669
8.8084 | -34.3377 | -15.0766
8.8146 | -34.3381 | -15.0863
8.8207 | -34.3384 | -15.0959
8.8268 | -34.3388 | -15.1053
8.8329 | -34.3391 | -15.115
8.839 | -34.3395 | -15.1247
8.8451 | -34.3399 | -15.1343
8.8512 | -34.3403 | -15.144
8.8573 | -34.3406 | -15.1537
8.8634 | -34.341 | -15.1632
8.8695 | -34.3414 | -15.1731
8.8756 | -34.3417 | -15.1828
8.8817 | -34.3421 | -15.1923
8.8878 | -34.3424 | -15.2022
8.8939 | -34.3428 | -15.2119
8.9 | -34.3432 | -15.2216
8.9061 | -34.3435 | -15.2313
8.9122 | -34.3439 | -15.2409
8.9183 | -34.3443 | -15.2506
8.9244 | -34.3447 | -15.2604
8.9305 | -34.3451 | -15.2699
8.9366 | -34.3455 | -15.2795
8.9427 | -34.3458 | -15.2893
8.9488 | -34.3462 | -15.2989
8.955 | -34.3466 | -15.3085
8.9611 | -34.347 | -15.3183
8.9672 | -34.3474 | -15.3278
8.9733 | -34.3477 | -15.3373
8.9794 | -34.3481 | -15.3468
8.9855 | -34.3485 | -15.3566
8.9916 | -34.3489 | -15.3664
8.9977 | -34.3493 | -15.376
9.0038 | -34.3497 | -15.3854
9.0099 | -34.3501 | -15.3951
9.016 | -34.3504 | -15.4048
9.0221 | -34.3508 | -15.4143
9.0282 | -34.3512 | -15.4239
9.0343 | -34.3516 | -15.4335
9.0404 | -34.352 | -15.4431
9.0465 | -34.3524 | -15.4529
9.0526 | -34.3528 | -15.4624
9.0587 | -34.3532 | -15.4722
9.0648 | -34.3536 | -15.4817
9.0709 | -34.3539 | -15.4913
9.077 | -34.3543 | -15.501
9.0831 | -34.3547 | -15.5107
9.0892 | -34.3551 | -15.5204
9.0953 | -34.3555 | -15.5301
9.1015 | -34.3559 | -15.5397
9.1076 | -34.3563 | -15.5491
9.1137 | -34.3567 | -15.5587
9.1198 | -34.3571 | -15.5683
9.1259 | -34.3575 | -15.5779
9.132 | -34.3578 | -15.5876
9.1381 | -34.3582 | -15.5971
9.1442 | -34.3587 | -15.6064
9.1503 | -34.359 | -15.6161
9.1564 | -34.3594 | -15.6256
9.1625 | -34.3598 | -15.6352
9.1686 | -34.3602 | -15.6447
9.1747 | -34.3606 | -15.6544
9.1808 | -34.361 | -15.6639
9.1869 | -34.3614 | -15.6734
9.193 | -34.3618 | -15.683
9.1991 | -34.3622 | -15.6927
9.2052 | -34.3626 | -15.7023
9.2113 | -34.363 | -15.7118
9.2174 | -34.3634 | -15.7213
9.2235 | -34.3638 | -15.7309
9.2296 | -34.3642 | -15.7404
9.2357 | -34.3646 | -15.7502
9.2419 | -34.365 | -15.7597
9.248 | -34.3655 | -15.7692
9.2541 | -34.3659 | -15.779
9.2602 | -34.3663 | -15.7887
9.2663 | -34.3667 | -15.798
9.2724 | -34.3671 | -15.8075
9.2785 | -34.3675 | -15.8169
9.2846 | -34.3679 | -15.8264
9.2907 | -34.3683 | -15.836
9.2968 | -34.3687 | -15.8456
9.3029 | -34.3691 | -15.8553
9.309 | -34.3695 | -15.8649
9.3151 | -34.3699 | -15.8744
9.3212 | -34.3703 | -15.8839
9.3273 | -34.3707 | -15.8933
9.3334 | -34.3711 | -15.9027
9.3395 | -34.3715 | -15.9122
9.3456 | -34.3719 | -15.9215
9.3517 | -34.3723 | -15.9311
9.3578 | -34.3727 | -15.9405
9.3639 | -34.3731 | -15.9499
9.37 | -34.3736 | -15.9596
9.3761 | -34.3739 | -15.9692
9.3822 | -34.3744 | -15.9788
9.3884 | -34.3748 | -15.9881
9.3945 | -34.3752 | -15.9977
9.4006 | -34.3756 | -16.0072
9.4067 | -34.376 | -16.0168
9.4128 | -34.3764 | -16.0264
9.4189 | -34.3768 | -16.0361
9.425 | -34.3773 | -16.0456
9.4311 | -34.3777 | -16.0551
9.4372 | -34.3781 | -16.0646
9.4433 | -34.3785 | -16.0742
9.4494 | -34.3789 | -16.0836
9.4555 | -34.3794 | -16.0931
9.4616 | -34.3798 | -16.1026
9.4677 | -34.3802 | -16.1121
9.4738 | -34.3806 | -16.1216
9.4799 | -34.381 | -16.1311
9.486 | -34.3814 | -16.1407
9.4921 | -34.3818 | -16.1502
9.4982 | -34.3823 | -16.1594
9.5043 | -34.3827 | -16.169
9.5104 | -34.3831 | -16.1783
9.5165 | -34.3835 | -16.1878
9.5226 | -34.3839 | -16.1975
9.5288 | -34.3843 | -16.2068
9.5349 | -34.3847 | -16.2164
9.541 | -34.3851 | -16.2257
9.5471 | -34.3855 | -16.2353
9.5532 | -34.3859 | -16.2448
9.5593 | -34.3864 | -16.2544
9.5654 | -34.3868 | -16.2638
9.5715 | -34.3872 | -16.2733
9.5776 | -34.3876 | -16.283
9.5837 | -34.388 | -16.2925
9.5898 | -34.3884 | -16.3021
9.5959 | -34.3888 | -16.3115
9.602 | -34.3892 | -16.321
9.6081 | -34.3896 | -16.3304
9.6142 | -34.3901 | -16.3398
9.6203 | -34.3905 | -16.3492
9.6264 | -34.3909 | -16.3586
9.6325 | -34.3913 | -16.3681
9.6386 | -34.3917 | -16.3774
9.6447 | -34.3921 | -16.3869
9.6508 | -34.3925 | -16.3964
9.6569 | -34.3929 | -16.4059
9.663 | -34.3934 | -16.4155
9.6691 | -34.3938 | -16.4249
9.6753 | -34.3942 | -16.4344
9.6814 | -34.3946 | -16.4438
9.6875 | -34.395 | -16.4532
9.6936 | -34.3954 | -16.4626
9.6997 | -34.3959 | -16.4723
9.7058 | -34.3963 | -16.4817
9.7119 | -34.3968 | -16.4912
9.718 | -34.3972 | -16.5008
9.7241 | -34.3976 | -16.5102
9.7302 | -34.398 | -16.5197
9.7363 | -34.3985 | -16.529
9.7424 | -34.3989 | -16.5385
9.7485 | -34.3993 | -16.548
9.7546 | -34.3997 | -16.5574
9.7607 | -34.4001 | -16.567
9.7668 | -34.4006 | -16.5766
9.7729 | -34.401 | -16.586
9.779 | -34.4014 | -16.5954
9.7851 | -34.4019 | -16.6048
9.7912 | -34.4023 | -16.6142
9.7973 | -34.4028 | -16.6237
9.8034 | -34.4032 | -16.6329
9.8095 | -34.4036 | -16.6424
9.8157 | -34.404 | -16.6518
9.8218 | -34.4044 | -16.6613
9.8279 | -34.4048 | -16.6706
9.834 | -34.4052 | -16.6802
9.8401 | -34.4056 | -16.6895
9.8462 | -34.4061 | -16.6987
9.8523 | -34.4065 | -16.7082
9.8584 | -34.4069 | -16.7177
9.8645 | -34.4073 | -16.7271
9.8706 | -34.4077 | -16.7365
9.8767 | -34.4081 | -16.7459
9.8828 | -34.4085 | -16.7551
9.8889 | -34.4089 | -16.7644
9.895 | -34.4094 | -16.7739
9.9011 | -34.4098 | -16.7833
9.9072 | -34.4102 | -16.7926
9.9133 | -34.4106 | -16.8021
9.9194 | -34.411 | -16.8113
9.9255 | -34.4115 | -16.8207
9.9316 | -34.4119 | -16.8302
9.9377 | -34.4123 | -16.8395
9.9438 | -34.4127 | -16.8491
9.9499 | -34.4132 | -16.8587
9.956 | -34.4136 | -16.8681
9.9622 | -34.414 | -16.8776
9.9683 | -34.4144 | -16.8871
9.9744 | -34.4148 | -16.8967
9.9805 | -34.4152 | -16.906
9.9866 | -34.4156 | -16.9151
9.9927 | -34.416 | -16.9244
9.9988 | -34.4164 | -16.934
10.0049 | -34.4169 | -16.9433
10.011 | -34.4173 | -16.9529
10.0171 | -34.4177 | -16.9621
10.0232 | -34.4182 | -16.9713
10.0293 | -34.4186 | -16.9807
10.0354 | -34.4191 | -16.9901
10.0415 | -34.4195 | -16.9995
10.0476 | -34.4199 | -17.0089
10.0537 | -34.4203 | -17.0185
10.0598 | -34.4207 | -17.028
10.0659 | -34.4211 | -17.0374
10.072 | -34.4215 | -17.0465
10.0781 | -34.422 | -17.0558
10.0842 | -34.4224 | -17.0652
10.0903 | -34.4228 | -17.0745
10.0964 | -34.4232 | -17.0838
10.1026 | -34.4236 | -17.0932
10.1087 | -34.424 | -17.1026
10.1148 | -34.4245 | -17.112
10.1209 | -34.4249 | -17.1214
10.127 | -34.4253 | -17.1306
10.1331 | -34.4257 | -17.14
10.1392 | -34.4261 | -17.1493
10.1453 | -34.4266 | -17.1586
10.1514 | -34.427 | -17.1682
10.1575 | -34.4274 | -17.1776
10.1636 | -34.4278 | -17.1868
10.1697 | -34.4282 | -17.1961
10.1758 | -34.4287 | -17.2054
10.1819 | -34.4291 | -17.2148
10.188 | -34.4295 | -17.2243
10.1941 | -34.4299 | -17.2336
10.2002 | -34.4303 | -17.243
10.2063 | -34.4307 | -17.2524
10.2124 | -34.4311 | -17.262
10.2185 | -34.4315 | -17.2715
10.2246 | -34.432 | -17.2808
10.2307 | -34.4324 | -17.2902
10.2368 | -34.4328 | -17.2995
10.2429 | -34.4332 | -17.309
10.2491 | -34.4336 | -17.3184
10.2552 | -34.434 | -17.3277
10.2613 | -34.4345 | -17.337
10.2674 | -34.4349 | -17.3464
10.2735 | -34.4353 | -17.3558
10.2796 | -34.4357 | -17.365
10.2857 | -34.4362 | -17.3744
10.2918 | -34.4366 | -17.3839
10.2979 | -34.437 | -17.3932
10.304 | -34.4374 | -17.4024
10.3101 | -34.4378 | -17.4119
10.3162 | -34.4382 | -17.4214
10.3223 | -34.4386 | -17.4306
10.3284 | -34.4391 | -17.4398
10.3345 | -34.4395 | -17.4492
10.3406 | -34.4399 | -17.4586
10.3467 | -34.4403 | -17.468
10.3528 | -34.4407 | -17.4773
10.3589 | -34.4411 | -17.4865
10.365 | -34.4415 | -17.496
10.3711 | -34.4419 | -17.5054
10.3772 | -34.4424 | -17.5148
10.3833 | -34.4428 | -17.5241
10.3895 | -34.4433 | -17.5334
10.3956 | -34.4437 | -17.5427
10.4017 | -34.4441 | -17.552
10.4078 | -34.4445 | -17.5613
10.4139 | -34.4449 | -17.5706
10.42 | -34.4454 | -17.5801
10.4261 | -34.4458 | -17.5894
10.4322 | -34.4463 | -17.5988
10.4383 | -34.4466 | -17.6081
10.4444 | -34.4471 | -17.6173
10.4505 | -34.4475 | -17.6266
10.4566 | -34.4479 | -17.6358
10.4627 | -34.4483 | -17.6451
10.4688 | -34.4487 | -17.6543
10.4749 | -34.4491 | -17.6636
10.481 | -34.4495 | -17.673
10.4871 | -34.45 | -17.6822
10.4932 | -34.4504 | -17.6914
10.4993 | -34.4508 | -17.7007
10.5054 | -34.4512 | -17.7101
10.5115 | -34.4516 | -17.7195
10.5176 | -34.452 | -17.7285
10.5237 | -34.4525 | -17.738
10.5298 | -34.4529 | -17.7474
10.536 | -34.4533 | -17.7566
10.5421 | -34.4537 | -17.7658
10.5482 | -34.4541 | -17.7752
10.5543 | -34.4546 | -17.7845
10.5604 | -34.455 | -17.7935
10.5665 | -34.4554 | -17.8028
10.5726 | -34.4558 | -17.812
10.5787 | -34.4562 | -17.8213
10.5848 | -34.4566 | -17.8302
10.5909 | -34.457 | -17.8396
10.597 | -34.4574 | -17.8489
10.6031 | -34.4578 | -17.8585
10.6092 | -34.4582 | -17.8678
10.6153 | -34.4586 | -17.8773
10.6214 | -34.4591 | -17.8863
10.6275 | -34.4595 | -17.8956
10.6336 | -34.4599 | -17.905
10.6397 | -34.4603 | -17.9143
10.6458 | -34.4607 | -17.9235
10.6519 | -34.4612 | -17.933
10.658 | -34.4616 | -17.9421
10.6641 | -34.462 | -17.9515
10.6702 | -34.4624 | -17.9608
10.6764 | -34.4628 | -17.9702
10.6825 | -34.4633 | -17.9794
10.6886 | -34.4636 | -17.9885
10.6947 | -34.4641 | -17.9977
10.7008 | -34.4645 | -18.007
10.7069 | -34.4649 | -18.0163
10.713 | -34.4653 | -18.0255
10.7191 | -34.4658 | -18.0347
10.7252 | -34.4662 | -18.0437
10.7313 | -34.4666 | -18.053
10.7374 | -34.467 | -18.0622
10.7435 | -34.4674 | -18.0715
10.7496 | -34.4678 | -18.0805
10.7557 | -34.4682 | -18.0899
10.7618 | -34.4687 | -18.099
10.7679 | -34.4691 | -18.1082
10.774 | -34.4695 | -18.1175
10.7801 | -34.47 | -18.1268
10.7862 | -34.4704 | -18.1358
10.7923 | -34.4708 | -18.1451
10.7984 | -34.4712 | -18.1541
10.8045 | -34.4716 | -18.1632
10.8106 | -34.4721 | -18.1724
10.8168 | -34.4725 | -18.1817
10.8229 | -34.4729 | -18.1909
10.829 | -34.4733 | -18.2002
10.8351 | -34.4737 | -18.2095
10.8412 | -34.4742 | -18.2187
10.8473 | -34.4747 | -18.2278
10.8534 | -34.4751 | -18.237
10.8595 | -34.4755 | -18.2462
10.8656 | -34.476 | -18.2552
10.8717 | -34.4764 | -18.2644
10.8778 | -34.4768 | -18.2737
10.8839 | -34.4772 | -18.283
10.89 | -34.4776 | -18.2923
10.8961 | -34.4781 | -18.3014
10.9022 | -34.4785 | -18.3105
10.9083 | -34.4789 | -18.3196
10.9144 | -34.4793 | -18.3287
10.9205 | -34.4797 | -18.3378
10.9266 | -34.4801 | -18.3471
10.9327 | -34.4806 | -18.3563
10.9388 | -34.481 | -18.3654
10.9449 | -34.4815 | -18.3746
10.951 | -34.4819 | -18.3838
10.9571 | -34.4823 | -18.3931
10.9633 | -34.4827 | -18.4024
10.9694 | -34.4832 | -18.4117
10.9755 | -34.4836 | -18.4209
10.9816 | -34.484 | -18.4301
10.9877 | -34.4844 | -18.4392
10.9938 | -34.4849 | -18.4484
10.9999 | -34.4853 | -18.4574
11.006 | -34.4858 | -18.4663
11.0121 | -34.4862 | -18.4755
11.0182 | -34.4866 | -18.4846
11.0243 | -34.487 | -18.4938
11.0304 | -34.4874 | -18.503
11.0365 | -34.4879 | -18.512
11.0426 | -34.4883 | -18.5212
11.0487 | -34.4888 | -18.5306
11.0548 | -34.4892 | -18.5396
11.0609 | -34.4896 | -18.5488
11.067 | -34.4901 | -18.5581
11.0731 | -34.4905 | -18.5671
11.0792 | -34.491 | -18.5762
11.0853 | -34.4914 | -18.5853
11.0914 | -34.4918 | -18.5946
11.0975 | -34.4923 | -18.6037
11.1037 | -34.4927 | -18.6128
11.1098 | -34.4931 | -18.622
11.1159 | -34.4936 | -18.6309
11.122 | -34.494 | -18.64
11.1281 | -34.4944 | -18.6492
11.1342 | -34.4948 | -18.6583
11.1403 | -34.4953 | -18.6675
11.1464 | -34.4957 | -18.6768
11.1525 | -34.4961 | -18.6858
11.1586 | -34.4966 | -18.695
11.1647 | -34.497 | -18.704
11.1708 | -34.4974 | -18.7131
11.1769 | -34.4978 | -18.7222
11.183 | -34.4983 | -18.7312
11.1891 | -34.4988 | -18.7402
11.1952 | -34.4992 | -18.7493
11.2013 | -34.4996 | -18.7585
11.2074 | -34.5 | -18.7676
11.2135 | -34.5005 | -18.777
11.2196 | -34.5009 | -18.7861
11.2257 | -34.5014 | -18.7954
11.2318 | -34.5018 | -18.8044
11.2379 | -34.5023 | -18.8137
11.244 | -34.5027 | -18.8229
11.2502 | -34.5032 | -18.832
11.2563 | -34.5037 | -18.8413
11.2624 | -34.5041 | -18.8503
11.2685 | -34.5045 | -18.8594
11.2746 | -34.5049 | -18.8684
11.2807 | -34.5054 | -18.8777
11.2868 | -34.5058 | -18.8869
11.2929 | -34.5063 | -18.8957
11.299 | -34.5067 | -18.9049
11.3051 | -34.5072 | -18.914
11.3112 | -34.5076 | -18.9233
11.3173 | -34.508 | -18.9322
11.3234 | -34.5085 | -18.9412
11.3295 | -34.509 | -18.9504
11.3356 | -34.5094 | -18.9595
11.3417 | -34.5098 | -18.9685
11.3478 | -34.5102 | -18.9778
11.3539 | -34.5107 | -18.9868
11.36 | -34.5111 | -18.9958
11.3661 | -34.5116 | -19.0048
11.3722 | -34.512 | -19.0137
11.3783 | -34.5125 | -19.0229
11.3844 | -34.5129 | -19.0321
11.3906 | -34.5134 | -19.0411
11.3967 | -34.5138 | -19.0501
11.4028 | -34.5143 | -19.0594
11.4089 | -34.5147 | -19.0684
11.415 | -34.5151 | -19.0773
11.4211 | -34.5156 | -19.0865
11.4272 | -34.516 | -19.0955
11.4333 | -34.5164 | -19.1043
11.4394 | -34.5169 | -19.1134
11.4455 | -34.5173 | -19.1224
11.4516 | -34.5177 | -19.1315
11.4577 | -34.5182 | -19.1406
11.4638 | -34.5187 | -19.1495
11.4699 | -34.5191 | -19.1586
11.476 | -34.5195 | -19.1679
11.4821 | -34.52 | -19.177
11.4882 | -34.5205 | -19.1862
11.4943 | -34.5209 | -19.1954
11.5004 | -34.5214 | -19.2043
11.5065 | -34.5218 | -19.2132
11.5126 | -34.5223 | -19.2222
11.5187 | -34.5227 | -19.2313
11.5248 | -34.5231 | -19.2405
11.5309 | -34.5236 | -19.2496
11.5371 | -34.524 | -19.2586
11.5432 | -34.5245 | -19.2677
11.5493 | -34.5249 | -19.2769
11.5554 | -34.5254 | -19.2859
11.5615 | -34.5258 | -19.2951
11.5676 | -34.5263 | -19.3042
11.5737 | -34.5268 | -19.3132
11.5798 | -34.5272 | -19.322
11.5859 | -34.5277 | -19.331
11.592 | -34.5281 | -19.3401
11.5981 | -34.5286 | -19.3491
11.6042 | -34.529 | -19.3581
11.6103 | -34.5295 | -19.367
11.6164 | -34.5299 | -19.3759
11.6225 | -34.5303 | -19.3848
11.6286 | -34.5308 | -19.3938
11.6347 | -34.5312 | -19.4029
11.6408 | -34.5317 | -19.4119
11.6469 | -34.5321 | -19.4209
11.653 | -34.5326 | -19.4298
11.6591 | -34.533 | -19.4388
11.6652 | -34.5335 | -19.4477
11.6713 | -34.5339 | -19.4566
11.6775 | -34.5344 | -19.4655
11.6836 | -34.5349 | -19.4745
11.6897 | -34.5353 | -19.4836
11.6958 | -34.5358 | -19.4926
11.7019 | -34.5362 | -19.5017
11.708 | -34.5367 | -19.5107
11.7141 | -34.5372 | -19.5195
11.7202 | -34.5376 | -19.5285
11.7263 | -34.538 | -19.5375
11.7324 | -34.5385 | -19.5467
11.7385 | -34.5389 | -19.5557
11.7446 | -34.5393 | -19.5647
11.7507 | -34.5398 | -19.5734
11.7568 | -34.5403 | -19.5824
11.7629 | -34.5407 | -19.5915
11.769 | -34.5412 | -19.6003
11.7751 | -34.5416 | -19.6093
11.7812 | -34.5421 | -19.6181
11.7873 | -34.5426 | -19.6271
11.7934 | -34.543 | -19.6362
11.7995 | -34.5435 | -19.6452
11.8056 | -34.544 | -19.654
11.8117 | -34.5444 | -19.6631
11.8178 | -34.5449 | -19.6719
11.824 | -34.5454 | -19.6808
11.8301 | -34.5458 | -19.6898
11.8362 | -34.5463 | -19.6986
11.8423 | -34.5467 | -19.7076
11.8484 | -34.5472 | -19.7166
11.8545 | -34.5477 | -19.7253
11.8606 | -34.5481 | -19.7342
11.8667 | -34.5485 | -19.7433
11.8728 | -34.549 | -19.7521
11.8789 | -34.5494 | -19.761
11.885 | -34.5499 | -19.7699
11.8911 | -34.5504 | -19.7789
11.8972 | -34.5508 | -19.7877
11.9033 | -34.5513 | -19.7965
11.9094 | -34.5517 | -19.8054
11.9155 | -34.5522 | -19.8142
11.9216 | -34.5527 | -19.823
11.9277 | -34.5531 | -19.8319
11.9338 | -34.5536 | -19.8408
11.9399 | -34.554 | -19.8498
11.946 | -34.5545 | -19.8587
11.9521 | -34.5549 | -19.8675
11.9582 | -34.5554 | -19.8763
11.9644 | -34.5558 | -19.8851
11.9705 | -34.5563 | -19.894
11.9766 | -34.5568 | -19.9029
11.9827 | -34.5573 | -19.9118
11.9888 | -34.5577 | -19.9206
11.9949 | -34.5582 | -19.9293
12.001 | -34.5587 | -19.9383
12.0071 | -34.5591 | -19.9471
12.0132 | -34.5596 | -19.9559
12.0193 | -34.56 | -19.9648
12.0254 | -34.5605 | -19.9737
12.0315 | -34.5609 | -19.9823
12.0376 | -34.5614 | -19.9911
12.0437 | -34.5618 | -19.9999
12.0498 | -34.5623 | -20.0088
12.0559 | -34.5627 | -20.0175
12.062 | -34.5632 | -20.026
12.0681 | -34.5637 | -20.035
12.0742 | -34.5641 | -20.0437
12.0803 | -34.5646 | -20.0527
12.0864 | -34.565 | -20.0617
12.0925 | -34.5655 | -20.0705
12.0986 | -34.566 | -20.0792
12.1047 | -34.5664 | -20.0881
12.1109 | -34.5669 | -20.0968
12.117 | -34.5674 | -20.1055
12.1231 | -34.5678 | -20.1144
12.1292 | -34.5683 | -20.1232
12.1353 | -34.5687 | -20.1319
12.1414 | -34.5692 | -20.1408
12.1475 | -34.5697 | -20.1497
12.1536 | -34.5701 | -20.1585
12.1597 | -34.5706 | -20.1673
12.1658 | -34.571 | -20.1761
12.1719 | -34.5715 | -20.1848
12.178 | -34.5719 | -20.1936
12.1841 | -34.5724 | -20.2025
12.1902 | -34.5729 | -20.2113
12.1963 | -34.5734 | -20.2202
12.2024 | -34.5739 | -20.2287
12.2085 | -34.5743 | -20.2375
12.2146 | -34.5748 | -20.2463
12.2207 | -34.5752 | -20.255
12.2268 | -34.5757 | -20.2638
12.2329 | -34.5762 | -20.2724
12.239 | -34.5767 | -20.2809
12.2451 | -34.5771 | -20.2897
12.2513 | -34.5776 | -20.2986
12.2574 | -34.578 | -20.3074
12.2635 | -34.5785 | -20.3161
12.2696 | -34.579 | -20.3248
12.2757 | -34.5795 | -20.3333
12.2818 | -34.5799 | -20.3421
12.2879 | -34.5804 | -20.3508
12.294 | -34.5809 | -20.3594
12.3001 | -34.5813 | -20.3681
12.3062 | -34.5818 | -20.3769
12.3123 | -34.5823 | -20.3853
12.3184 | -34.5827 | -20.394
12.3245 | -34.5832 | -20.4026
12.3306 | -34.5837 | -20.4113
12.3367 | -34.5841 | -20.42
12.3428 | -34.5846 | -20.4287
12.3489 | -34.585 | -20.4371
12.355 | -34.5855 | -20.446
12.3611 | -34.586 | -20.4548
12.3672 | -34.5864 | -20.4635
12.3733 | -34.5869 | -20.4722
12.3794 | -34.5874 | -20.4806
12.3855 | -34.5878 | -20.4892
12.3916 | -34.5883 | -20.4979
12.3978 | -34.5888 | -20.5068
12.4039 | -34.5892 | -20.5155
12.41 | -34.5897 | -20.5242
12.4161 | -34.5902 | -20.5325
12.4222 | -34.5906 | -20.5411
12.4283 | -34.5911 | -20.5498
12.4344 | -34.5915 | -20.5585
12.4405 | -34.592 | -20.5673
12.4466 | -34.5924 | -20.5758
12.4527 | -34.5929 | -20.5844
12.4588 | -34.5934 | -20.5931
12.4649 | -34.5938 | -20.6018
12.471 | -34.5943 | -20.6106
12.4771 | -34.5948 | -20.6192
12.4832 | -34.5952 | -20.6278
12.4893 | -34.5957 | -20.6363
12.4954 | -34.5961 | -20.6449
12.5015 | -34.5966 | -20.6535
12.5076 | -34.5971 | -20.6622
12.5137 | -34.5976 | -20.6709
12.5198 | -34.598 | -20.6793
12.5259 | -34.5985 | -20.688
12.532 | -34.5989 | -20.6967
12.5382 | -34.5994 | -20.7053
12.5443 | -34.5999 | -20.714
12.5504 | -34.6003 | -20.7226
12.5565 | -34.6008 | -20.7311
12.5626 | -34.6013 | -20.7398
12.5687 | -34.6017 | -20.7485
12.5748 | -34.6022 | -20.7571
12.5809 | -34.6027 | -20.7657
12.587 | -34.6031 | -20.7744
12.5931 | -34.6036 | -20.7828
12.5992 | -34.6041 | -20.7913
12.6053 | -34.6046 | -20.7999
12.6114 | -34.605 | -20.8085
12.6175 | -34.6055 | -20.8172
12.6236 | -34.6059 | -20.8257
12.6297 | -34.6064 | -20.8342
12.6358 | -34.6069 | -20.8426
12.6419 | -34.6073 | -20.8512
12.648 | -34.6078 | -20.8599
12.6541 | -34.6082 | -20.8684
12.6602 | -34.6087 | -20.8766
12.6663 | -34.6092 | -20.8851
12.6724 | -34.6097 | -20.8936
12.6785 | -34.6101 | -20.9022
12.6847 | -34.6106 | -20.9107
12.6908 | -34.611 | -20.9193
12.6969 | -34.6115 | -20.9276
12.703 | -34.612 | -20.9362
12.7091 | -34.6124 | -20.9449
12.7152 | -34.6129 | -20.9533
12.7213 | -34.6133 | -20.9619
12.7274 | -34.6138 | -20.9705
12.7335 | -34.6143 | -20.9787
12.7396 | -34.6147 | -20.9872
12.7457 | -34.6152 | -20.9959
12.7518 | -34.6156 | -21.0044
12.7579 | -34.6161 | -21.013
12.764 | -34.6166 | -21.0216
12.7701 | -34.617 | -21.0301
12.7762 | -34.6175 | -21.0385
12.7823 | -34.618 | -21.0471
12.7884 | -34.6184 | -21.0557
12.7945 | -34.6189 | -21.0643
12.8006 | -34.6194 | -21.0728
12.8067 | -34.6198 | -21.0813
12.8128 | -34.6203 | -21.0896
12.8189 | -34.6207 | -21.0981
12.8251 | -34.6212 | -21.1066
12.8312 | -34.6216 | -21.1153
12.8373 | -34.6221 | -21.1237
12.8434 | -34.6226 | -21.1322
12.8495 | -34.623 | -21.1405
12.8556 | -34.6235 | -21.1489
12.8617 | -34.6239 | -21.1575
12.8678 | -34.6244 | -21.1661
12.8739 | -34.6249 | -21.1745
12.88 | -34.6253 | -21.1829
12.8861 | -34.6257 | -21.1913
12.8922 | -34.6262 | -21.1998
12.8983 | -34.6267 | -21.2084
12.9044 | -34.6271 | -21.2166
12.9105 | -34.6276 | -21.225
12.9166 | -34.6281 | -21.2335
12.9227 | -34.6285 | -21.2419
12.9288 | -34.629 | -21.2504
12.9349 | -34.6294 | -21.259
12.941 | -34.6299 | -21.2674
12.9471 | -34.6304 | -21.2759
12.9532 | -34.6309 | -21.2844
12.9593 | -34.6313 | -21.293
12.9654 | -34.6318 | -21.3015
12.9716 | -34.6323 | -21.3101
12.9777 | -34.6327 | -21.3185
12.9838 | -34.6332 | -21.327
12.9899 | -34.6337 | -21.3355
12.996 | -34.6341 | -21.3439
13.0021 | -34.6346 | -21.3523
13.0082 | -34.6351 | -21.3606
13.0143 | -34.6355 | -21.3688
13.0204 | -34.636 | -21.3771
13.0265 | -34.6364 | -21.3855
13.0326 | -34.6369 | -21.394
13.0387 | -34.6374 | -21.4025
13.0448 | -34.6378 | -21.4109
13.0509 | -34.6383 | -21.4193
13.057 | -34.6388 | -21.4276
13.0631 | -34.6393 | -21.436
13.0692 | -34.6397 | -21.4444
13.0753 | -34.6402 | -21.453
13.0814 | -34.6406 | -21.4617
13.0875 | -34.6411 | -21.4698
13.0936 | -34.6415 | -21.4784
13.0997 | -34.642 | -21.4867
13.1058 | -34.6425 | -21.4951
13.112 | -34.6429 | -21.5035
13.1181 | -34.6434 | -21.5118
13.1242 | -34.6439 | -21.5202
13.1303 | -34.6443 | -21.5286
13.1364 | -34.6448 | -21.537
13.1425 | -34.6452 | -21.5454
13.1486 | -34.6457 | -21.5538
13.1547 | -34.6461 | -21.5622
13.1608 | -34.6466 | -21.5706
13.1669 | -34.647 | -21.579
13.173 | -34.6475 | -21.5874
13.1791 | -34.648 | -21.5957
13.1852 | -34.6484 | -21.6042
13.1913 | -34.6489 | -21.6124
13.1974 | -34.6493 | -21.6209
13.2035 | -34.6497 | -21.6293
13.2096 | -34.6502 | -21.6376
13.2157 | -34.6507 | -21.646
13.2218 | -34.6512 | -21.6542
13.2279 | -34.6516 | -21.6624
13.234 | -34.6521 | -21.6707
13.2401 | -34.6526 | -21.679
13.2462 | -34.653 | -21.6873
13.2524 | -34.6535 | -21.6957
13.2585 | -34.6539 | -21.7039
13.2646 | -34.6544 | -21.7123
13.2707 | -34.6549 | -21.7207
13.2768 | -34.6554 | -21.7291
13.2829 | -34.6558 | -21.7376
13.289 | -34.6563 | -21.7459
13.2951 | -34.6567 | -21.7541
13.3012 | -34.6572 | -21.7625
13.3073 | -34.6577 | -21.7709
13.3134 | -34.6581 | -21.7793
13.3195 | -34.6586 | -21.7875
13.3256 | -34.6591 | -21.7958
13.3317 | -34.6595 | -21.804
13.3378 | -34.66 | -21.8124
13.3439 | -34.6605 | -21.8208
13.35 | -34.6609 | -21.8292
13.3561 | -34.6614 | -21.8374
13.3622 | -34.6618 | -21.8457
13.3683 | -34.6623 | -21.854
13.3744 | -34.6628 | -21.8622
13.3805 | -34.6632 | -21.8706
13.3866 | -34.6637 | -21.8789
13.3927 | -34.6641 | -21.8873
13.3989 | -34.6646 | -21.8955
13.405 | -34.665 | -21.9038
13.4111 | -34.6655 | -21.912
13.4172 | -34.666 | -21.9203
13.4233 | -34.6664 | -21.9287
13.4294 | -34.6669 | -21.9372
13.4355 | -34.6673 | -21.9452
13.4416 | -34.6678 | -21.9537
13.4477 | -34.6683 | -21.9619
13.4538 | -34.6687 | -21.9703
13.4599 | -34.6692 | -21.9787
13.466 | -34.6696 | -21.9871
13.4721 | -34.6701 | -21.9953
13.4782 | -34.6706 | -22.0037
13.4843 | -34.6711 | -22.0121
13.4904 | -34.6715 | -22.0203
13.4965 | -34.672 | -22.0288
13.5026 | -34.6725 | -22.0372
13.5087 | -34.6729 | -22.0452
13.5148 | -34.6734 | -22.0536
13.5209 | -34.6739 | -22.0621
13.527 | -34.6743 | -22.0704
13.5331 | -34.6748 | -22.0785
13.5393 | -34.6752 | -22.0865
13.5454 | -34.6757 | -22.0947
13.5515 | -34.6762 | -22.1031
13.5576 | -34.6766 | -22.1114
13.5637 | -34.6771 | -22.1198
13.5698 | -34.6776 | -22.128
13.5759 | -34.678 | -22.136
13.582 | -34.6785 | -22.1442
13.5881 | -34.679 | -22.1525
13.5942 | -34.6794 | -22.1609
13.6003 | -34.6799 | -22.1691
13.6064 | -34.6803 | -22.1775
13.6125 | -34.6808 | -22.1859
13.6186 | -34.6813 | -22.1942
13.6247 | -34.6818 | -22.2025
13.6308 | -34.6822 | -22.2109
13.6369 | -34.6827 | -22.2193
13.643 | -34.6832 | -22.2276
13.6491 | -34.6836 | -22.2359
13.6552 | -34.6841 | -22.2442
13.6613 | -34.6846 | -22.2524
13.6674 | -34.685 | -22.2607
13.6735 | -34.6855 | -22.2691
13.6796 | -34.6859 | -22.2774
13.6858 | -34.6864 | -22.2855
13.6919 | -34.6868 | -22.2938
13.698 | -34.6873 | -22.3022
13.7041 | -34.6878 | -22.3106
13.7102 | -34.6882 | -22.3189
13.7163 | -34.6887 | -22.327
13.7224 | -34.6892 | -22.3352
13.7285 | -34.6896 | -22.3434
13.7346 | -34.6901 | -22.3517
13.7407 | -34.6905 | -22.36
13.7468 | -34.691 | -22.3683
13.7529 | -34.6915 | -22.3765
13.759 | -34.6919 | -22.3847
13.7651 | -34.6924 | -22.393
13.7712 | -34.6928 | -22.4013
13.7773 | -34.6933 | -22.4095
13.7834 | -34.6938 | -22.4177
13.7895 | -34.6942 | -22.4259
13.7956 | -34.6947 | -22.4343
13.8017 | -34.6951 | -22.4424
13.8078 | -34.6956 | -22.4507
13.8139 | -34.696 | -22.4589
13.82 | -34.6965 | -22.4671
13.8262 | -34.697 | -22.4754
13.8323 | -34.6974 | -22.4837
13.8384 | -34.6979 | -22.4921
13.8445 | -34.6984 | -22.5004
13.8506 | -34.6988 | -22.5086
13.8567 | -34.6993 | -22.5169
13.8628 | -34.6998 | -22.5252
13.8689 | -34.7003 | -22.5334
13.875 | -34.7007 | -22.5417
13.8811 | -34.7012 | -22.5499
13.8872 | -34.7016 | -22.5582
13.8933 | -34.7021 | -22.5663
13.8994 | -34.7025 | -22.5746
13.9055 | -34.703 | -22.5828
13.9116 | -34.7035 | -22.5911
13.9177 | -34.7039 | -22.5996
13.9238 | -34.7044 | -22.6079
13.9299 | -34.7049 | -22.616
13.936 | -34.7053 | -22.6245
13.9421 | -34.7058 | -22.6327
13.9482 | -34.7063 | -22.6409
13.9543 | -34.7067 | -22.6493
13.9604 | -34.7072 | -22.6576
13.9665 | -34.7077 | -22.6658
13.9727 | -34.7082 | -22.6742
13.9788 | -34.7086 | -22.6824
13.9849 | -34.7091 | -22.6908
13.991 | -34.7096 | -22.6992
13.9971 | -34.71 | -22.7073
14.0032 | -34.7105 | -22.7156
14.0093 | -34.7109 | -22.7238
14.0154 | -34.7114 | -22.732
14.0215 | -34.7118 | -22.7401
14.0276 | -34.7123 | -22.7484
14.0337 | -34.7128 | -22.7566
14.0398 | -34.7133 | -22.7649
14.0459 | -34.7138 | -22.7731
14.052 | -34.7142 | -22.7815
14.0581 | -34.7147 | -22.7897
14.0642 | -34.7151 | -22.798
14.0703 | -34.7156 | -22.8062
14.0764 | -34.7161 | -22.8142
14.0825 | -34.7165 | -22.8224
14.0886 | -34.717 | -22.8306
14.0947 | -34.7175 | -22.8389
14.1008 | -34.7179 | -22.8471
14.1069 | -34.7184 | -22.8553
14.1131 | -34.7188 | -22.8636
14.1192 | -34.7193 | -22.8718
14.1253 | -34.7198 | -22.8799
14.1314 | -34.7202 | -22.8881
14.1375 | -34.7207 | -22.8962
14.1436 | -34.7212 | -22.9044
14.1497 | -34.7217 | -22.9126
14.1558 | -34.7221 | -22.9209
14.1619 | -34.7226 | -22.929
14.168 | -34.7231 | -22.9372
14.1741 | -34.7236 | -22.9454
14.1802 | -34.724 | -22.9535
14.1863 | -34.7245 | -22.9618
14.1924 | -34.725 | -22.9701
14.1985 | -34.7255 | -22.9784
14.2046 | -34.726 | -22.9867
14.2107 | -34.7264 | -22.9946
14.2168 | -34.7269 | -23.0029
14.2229 | -34.7273 | -23.0111
14.229 | -34.7278 | -23.0193
14.2351 | -34.7283 | -23.0276
14.2412 | -34.7287 | -23.0359
14.2473 | -34.7292 | -23.044
14.2534 | -34.7297 | -23.0521
14.2596 | -34.7301 | -23.0602
14.2657 | -34.7305 | -23.0684
14.2718 | -34.731 | -23.0767
14.2779 | -34.7315 | -23.085
14.284 | -34.732 | -23.0932
14.2901 | -34.7324 | -23.1014
14.2962 | -34.7329 | -23.1095
14.3023 | -34.7334 | -23.1178
14.3084 | -34.7338 | -23.126
14.3145 | -34.7343 | -23.1342
14.3206 | -34.7348 | -23.1427
14.3267 | -34.7352 | -23.1508
14.3328 | -34.7357 | -23.1589
14.3389 | -34.7362 | -23.1671
14.345 | -34.7367 | -23.1752
14.3511 | -34.7371 | -23.1834
14.3572 | -34.7376 | -23.1917
14.3633 | -34.7381 | -23.2
14.3694 | -34.7385 | -23.2082
14.3755 | -34.739 | -23.2164
14.3816 | -34.7395 | -23.2247
14.3877 | -34.74 | -23.2327
14.3938 | -34.7405 | -23.2411
14.4 | -34.7409 | -23.2492
14.4061 | -34.7414 | -23.2573
14.4122 | -34.7419 | -23.2655
14.4183 | -34.7423 | -23.2736
14.4244 | -34.7428 | -23.2817
14.4305 | -34.7433 | -23.2898
14.4366 | -34.7438 | -23.2979
14.4427 | -34.7443 | -23.3063
14.4488 | -34.7447 | -23.3145
14.4549 | -34.7452 | -23.3227
14.461 | -34.7457 | -23.3309
14.4671 | -34.7462 | -23.3391
14.4732 | -34.7466 | -23.3473
14.4793 | -34.7471 | -23.3554
14.4854 | -34.7476 | -23.3636
14.4915 | -34.7481 | -23.3716
14.4976 | -34.7486 | -23.3796
14.5037 | -34.7491 | -23.3878
14.5098 | -34.7495 | -23.396
14.5159 | -34.75 | -23.4041
14.522 | -34.7504 | -23.4121
14.5281 | -34.7509 | -23.4202
14.5342 | -34.7514 | -23.4284
14.5403 | -34.7519 | -23.4365
14.5465 | -34.7523 | -23.4446
14.5526 | -34.7528 | -23.4528
14.5587 | -34.7533 | -23.4609
14.5648 | -34.7538 | -23.469
14.5709 | -34.7543 | -23.4771
14.577 | -34.7547 | -23.4853
14.5831 | -34.7552 | -23.4935
14.5892 | -34.7557 | -23.5016
14.5953 | -34.7561 | -23.5097
14.6014 | -34.7566 | -23.518
14.6075 | -34.7571 | -23.5261
14.6136 | -34.7575 | -23.5345
14.6197 | -34.758 | -23.5428
14.6258 | -34.7584 | -23.5509
14.6319 | -34.7589 | -23.559
14.638 | -34.7594 | -23.5671
14.6441 | -34.7598 | -23.5754
14.6502 | -34.7603 | -23.5836
14.6563 | -34.7608 | -23.5918
14.6624 | -34.7613 | -23.5999
14.6685 | -34.7618 | -23.6078
14.6746 | -34.7623 | -23.616
14.6807 | -34.7627 | -23.6241
14.6869 | -34.7632 | -23.6322
14.693 | -34.7637 | -23.6402
14.6991 | -34.7642 | -23.6483
14.7052 | -34.7647 | -23.6564
14.7113 | -34.7652 | -23.6645
14.7174 | -34.7656 | -23.6727
14.7235 | -34.7661 | -23.6808
14.7296 | -34.7666 | -23.6889
14.7357 | -34.7671 | -23.6969
14.7418 | -34.7675 | -23.705
14.7479 | -34.768 | -23.713
14.754 | -34.7685 | -23.7212
14.7601 | -34.769 | -23.7293
14.7662 | -34.7694 | -23.7373
14.7723 | -34.7699 | -23.7453
14.7784 | -34.7704 | -23.7534
14.7845 | -34.7709 | -23.7616
14.7906 | -34.7714 | -23.7698
14.7967 | -34.7718 | -23.7779
14.8028 | -34.7723 | -23.786
14.8089 | -34.7728 | -23.7941
14.815 | -34.7733 | -23.8022
14.8211 | -34.7737 | -23.8103
14.8272 | -34.7742 | -23.8183
14.8334 | -34.7747 | -23.8265
14.8395 | -34.7752 | -23.8345
14.8456 | -34.7757 | -23.8425
14.8517 | -34.7761 | -23.8509
14.8578 | -34.7766 | -23.8591
14.8639 | -34.7771 | -23.8673
14.87 | -34.7776 | -23.8754
14.8761 | -34.7781 | -23.8834
14.8822 | -34.7785 | -23.8916
14.8883 | -34.779 | -23.8998
14.8944 | -34.7795 | -23.9079
14.9005 | -34.78 | -23.916
14.9066 | -34.7804 | -23.9241
14.9127 | -34.7809 | -23.9322
14.9188 | -34.7814 | -23.9402
14.9249 | -34.7818 | -23.9484
14.931 | -34.7823 | -23.9566
14.9371 | -34.7828 | -23.9645
14.9432 | -34.7833 | -23.9727
14.9493 | -34.7838 | -23.9808
14.9554 | -34.7843 | -23.9889
14.9615 | -34.7847 | -23.9971
14.9676 | -34.7852 | -24.0051
14.9738 | -34.7857 | -24.0133
14.9799 | -34.7861 | -24.0215
14.986 | -34.7866 | -24.0295
14.9921 | -34.7871 | -24.0376
14.9982 | -34.7876 | -24.0458
15.0043 | -34.788 | -24.0537
15.0104 | -34.7885 | -24.062
15.0165 | -34.789 | -24.0702
15.0226 | -34.7895 | -24.078
15.0287 | -34.79 | -24.0863
15.0348 | -34.7905 | -24.0944
15.0409 | -34.7909 | -24.1026
15.047 | -34.7914 | -24.1108
15.0531 | -34.792 | -24.1188
15.0592 | -34.7924 | -24.127
15.0653 | -34.7929 | -24.1351
15.0714 | -34.7934 | -24.143
15.0775 | -34.7939 | -24.1509
15.0836 | -34.7944 | -24.159
15.0897 | -34.7949 | -24.1669
15.0958 | -34.7954 | -24.175
15.1019 | -34.7958 | -24.1831
15.108 | -34.7963 | -24.1911
15.1141 | -34.7968 | -24.1992
15.1203 | -34.7973 | -24.2074
15.1264 | -34.7978 | -24.2153
15.1325 | -34.7982 | -24.2233
15.1386 | -34.7987 | -24.2313
15.1447 | -34.7992 | -24.2394
15.1508 | -34.7997 | -24.2475
15.1569 | -34.8002 | -24.2555
15.163 | -34.8007 | -24.2633
15.1691 | -34.8011 | -24.2713
15.1752 | -34.8016 | -24.2792
15.1813 | -34.8021 | -24.2873
15.1874 | -34.8026 | -24.2954
15.1935 | -34.8031 | -24.3034
15.1996 | -34.8036 | -24.3114
15.2057 | -34.8041 | -24.3194
15.2118 | -34.8046 | -24.3275
15.2179 | -34.805 | -24.3356
15.224 | -34.8055 | -24.3437
15.2301 | -34.806 | -24.3515
15.2362 | -34.8065 | -24.3595
15.2423 | -34.807 | -24.3676
15.2484 | -34.8074 | -24.3757
15.2545 | -34.8079 | -24.3837
15.2607 | -34.8084 | -24.3919
15.2668 | -34.8089 | -24.3997
15.2729 | -34.8094 | -24.4077
15.279 | -34.8099 | -24.4159
15.2851 | -34.8104 | -24.4239
15.2912 | -34.8109 | -24.4319
15.2973 | -34.8114 | -24.44
15.3034 | -34.8118 | -24.4479
15.3095 | -34.8123 | -24.4559
15.3156 | -34.8128 | -24.4639
15.3217 | -34.8133 | -24.472
15.3278 | -34.8138 | -24.4801
15.3339 | -34.8143 | -24.488
15.34 | -34.8148 | -24.496
15.3461 | -34.8152 | -24.5041
15.3522 | -34.8157 | -24.512
15.3583 | -34.8162 | -24.5202
15.3644 | -34.8167 | -24.5282
15.3705 | -34.8171 | -24.5359
15.3766 | -34.8176 | -24.5439
15.3827 | -34.8181 | -24.5519
15.3888 | -34.8186 | -24.5597
15.3949 | -34.8191 | -24.5677
15.401 | -34.8196 | -24.5756
15.4072 | -34.8201 | -24.5834
15.4133 | -34.8205 | -24.5914
15.4194 | -34.821 | -24.5994
15.4255 | -34.8215 | -24.6073
15.4316 | -34.8219 | -24.6153
15.4377 | -34.8224 | -24.6235
15.4438 | -34.8229 | -24.6314
15.4499 | -34.8234 | -24.6393
15.456 | -34.8239 | -24.6472
15.4621 | -34.8244 | -24.6551
15.4682 | -34.8249 | -24.6631
15.4743 | -34.8254 | -24.6708
15.4804 | -34.8258 | -24.6787
15.4865 | -34.8263 | -24.6868
15.4926 | -34.8268 | -24.6948
15.4987 | -34.8273 | -24.7027
15.5048 | -34.8278 | -24.7108
15.5109 | -34.8283 | -24.7187
15.517 | -34.8288 | -24.7267
15.5231 | -34.8293 | -24.7346
15.5292 | -34.8298 | -24.7425
15.5353 | -34.8303 | -24.7504
15.5414 | -34.8308 | -24.7583
15.5476 | -34.8313 | -24.7661
15.5537 | -34.8317 | -24.774
15.5598 | -34.8322 | -24.782
15.5659 | -34.8327 | -24.7899
15.572 | -34.8332 | -24.7979
15.5781 | -34.8336 | -24.8057
15.5842 | -34.8341 | -24.8135
15.5903 | -34.8346 | -24.8213
15.5964 | -34.8351 | -24.8292
15.6025 | -34.8356 | -24.837
15.6086 | -34.8361 | -24.8449
15.6147 | -34.8366 | -24.8529
15.6208 | -34.8371 | -24.8607
15.6269 | -34.8376 | -24.8687
15.633 | -34.8381 | -24.8767
15.6391 | -34.8386 | -24.8847
15.6452 | -34.839 | -24.8928
15.6513 | -34.8395 | -24.9005
15.6574 | -34.84 | -24.9086
15.6635 | -34.8405 | -24.9166
15.6696 | -34.841 | -24.9244
15.6757 | -34.8415 | -24.9324
15.6818 | -34.842 | -24.9404
15.688 | -34.8425 | -24.9482
15.6941 | -34.8429 | -24.956
15.7002 | -34.8434 | -24.9639
15.7063 | -34.8439 | -24.9717
15.7124 | -34.8444 | -24.9797
15.7185 | -34.8449 | -24.9877
15.7246 | -34.8454 | -24.9954
15.7307 | -34.8459 | -25.0032
15.7368 | -34.8464 | -25.0111
15.7429 | -34.8469 | -25.019
15.749 | -34.8474 | -25.0269
15.7551 | -34.8479 | -25.035
15.7612 | -34.8484 | -25.0427
15.7673 | -34.8489 | -25.0505
15.7734 | -34.8493 | -25.0584
15.7795 | -34.8498 | -25.0663
15.7856 | -34.8503 | -25.0741
15.7917 | -34.8508 | -25.0817
15.7978 | -34.8513 | -25.0897
15.8039 | -34.8518 | -25.0977
15.81 | -34.8522 | -25.1055
15.8161 | -34.8527 | -25.1134
15.8222 | -34.8532 | -25.1214
15.8283 | -34.8537 | -25.129
15.8345 | -34.8542 | -25.1368
15.8406 | -34.8547 | -25.1447
15.8467 | -34.8552 | -25.1527
15.8528 | -34.8557 | -25.1605
15.8589 | -34.8562 | -25.1684
15.865 | -34.8567 | -25.176
15.8711 | -34.8572 | -25.184
15.8772 | -34.8577 | -25.1917
15.8833 | -34.8582 | -25.1997
15.8894 | -34.8587 | -25.2076
15.8955 | -34.8591 | -25.2153
15.9016 | -34.8596 | -25.223
15.9077 | -34.8601 | -25.2309
15.9138 | -34.8606 | -25.2389
15.9199 | -34.8611 | -25.2466
15.926 | -34.8616 | -25.2544
15.9321 | -34.8621 | -25.2621
15.9382 | -34.8625 | -25.2698
15.9443 | -34.863 | -25.2777
15.9504 | -34.8635 | -25.2855
15.9565 | -34.864 | -25.2932
15.9626 | -34.8645 | -25.3011
15.9687 | -34.865 | -25.3086
15.9749 | -34.8655 | -25.3162
15.981 | -34.866 | -25.3241
15.9871 | -34.8665 | -25.332
15.9932 | -34.8669 | -25.3396
15.9993 | -34.8674 | -25.3474
16.0054 | -34.8679 | -25.3548
16.0115 | -34.8685 | -25.3626
16.0176 | -34.8689 | -25.3706
16.0237 | -34.8694 | -25.3785
16.0298 | -34.8699 | -25.3862
16.0359 | -34.8704 | -25.3941
16.042 | -34.8709 | -25.4017
16.0481 | -34.8714 | -25.4095
16.0542 | -34.8719 | -25.4174
16.0603 | -34.8724 | -25.425
16.0664 | -34.8728 | -25.4329
16.0725 | -34.8734 | -25.4407
16.0786 | -34.8738 | -25.4484
16.0847 | -34.8743 | -25.4564
16.0908 | -34.8748 | -25.4642
16.0969 | -34.8753 | -25.472
16.103 | -34.8757 | -25.4798
16.1091 | -34.8763 | -25.4874
16.1152 | -34.8767 | -25.4953
16.1214 | -34.8772 | -25.5031
16.1275 | -34.8777 | -25.5109
16.1336 | -34.8782 | -25.5188
16.1397 | -34.8787 | -25.5265
16.1458 | -34.8792 | -25.5342
16.1519 | -34.8797 | -25.542
16.158 | -34.8802 | -25.5498
16.1641 | -34.8806 | -25.5576
16.1702 | -34.8811 | -25.5654
16.1763 | -34.8816 | -25.5733
16.1824 | -34.8821 | -25.5807
16.1885 | -34.8826 | -25.5886
16.1946 | -34.8831 | -25.5962
16.2007 | -34.8836 | -25.6039
16.2068 | -34.8841 | -25.6116
16.2129 | -34.8846 | -25.6193
16.219 | -34.8851 | -25.6267
16.2251 | -34.8856 | -25.6345
16.2312 | -34.8861 | -25.6422
16.2373 | -34.8866 | -25.6499
16.2434 | -34.8871 | -25.6577
16.2495 | -34.8876 | -25.6652
16.2556 | -34.8881 | -25.6731
16.2618 | -34.8886 | -25.6808
16.2679 | -34.8891 | -25.6885
16.274 | -34.8896 | -25.6962
16.2801 | -34.89 | -25.7038
16.2862 | -34.8905 | -25.7116
16.2923 | -34.891 | -25.7193
16.2984 | -34.8915 | -25.727
16.3045 | -34.892 | -25.7348
16.3106 | -34.8925 | -25.7426
16.3167 | -34.893 | -25.7503
16.3228 | -34.8935 | -25.7578
16.3289 | -34.8939 | -25.7655
16.335 | -34.8944 | -25.7733
16.3411 | -34.8949 | -25.781
16.3472 | -34.8954 | -25.7887
16.3533 | -34.896 | -25.7965
16.3594 | -34.8965 | -25.8041
16.3655 | -34.8969 | -25.812
16.3716 | -34.8974 | -25.8195
16.3777 | -34.8979 | -25.8272
16.3838 | -34.8984 | -25.8347
16.3899 | -34.8989 | -25.8423
16.396 | -34.8994 | -25.8499
16.4021 | -34.8999 | -25.8576
16.4083 | -34.9003 | -25.8651
16.4144 | -34.9008 | -25.873
16.4205 | -34.9013 | -25.8806
16.4266 | -34.9018 | -25.888
16.4327 | -34.9023 | -25.8959
16.4388 | -34.9028 | -25.9034
16.4449 | -34.9033 | -25.911
16.451 | -34.9038 | -25.9187
16.4571 | -34.9043 | -25.9265
16.4632 | -34.9048 | -25.9339
16.4693 | -34.9053 | -25.9415
16.4754 | -34.9058 | -25.9491
16.4815 | -34.9062 | -25.9569
16.4876 | -34.9067 | -25.9645
16.4937 | -34.9072 | -25.9721
16.4998 | -34.9077 | -25.9795
16.5059 | -34.9082 | -25.9871
16.512 | -34.9087 | -25.9949
16.5181 | -34.9092 | -26.0029
16.5242 | -34.9097 | -26.0107
16.5303 | -34.9102 | -26.0184
16.5364 | -34.9107 | -26.0259
16.5425 | -34.9112 | -26.0337
16.5487 | -34.9117 | -26.0411
16.5548 | -34.9122 | -26.0488
16.5609 | -34.9126 | -26.0565
16.567 | -34.9132 | -26.064
16.5731 | -34.9137 | -26.0719
16.5792 | -34.9141 | -26.0796
16.5853 | -34.9146 | -26.0873
16.5914 | -34.9151 | -26.0951
16.5975 | -34.9155 | -26.1029
16.6036 | -34.916 | -26.1106
16.6097 | -34.9165 | -26.1183
16.6158 | -34.917 | -26.1259
16.6219 | -34.9175 | -26.1336
16.628 | -34.918 | -26.1411
16.6341 | -34.9184 | -26.1488
16.6402 | -34.919 | -26.1563
16.6463 | -34.9194 | -26.164
16.6524 | -34.9199 | -26.1717
16.6585 | -34.9204 | -26.1794
16.6646 | -34.9209 | -26.187
16.6707 | -34.9214 | -26.1945
16.6768 | -34.9219 | -26.2022
16.6829 | -34.9224 | -26.2097
16.689 | -34.9229 | -26.2174
16.6952 | -34.9234 | -26.225
16.7013 | -34.9239 | -26.2325
16.7074 | -34.9243 | -26.24
16.7135 | -34.9248 | -26.2476
16.7196 | -34.9253 | -26.2552
16.7257 | -34.9258 | -26.2629
16.7318 | -34.9263 | -26.2706
16.7379 | -34.9268 | -26.2783
16.744 | -34.9272 | -26.2857
16.7501 | -34.9277 | -26.2933
16.7562 | -34.9282 | -26.301
16.7623 | -34.9287 | -26.3086
16.7684 | -34.9292 | -26.3162
16.7745 | -34.9296 | -26.324
16.7806 | -34.9302 | -26.3313
16.7867 | -34.9307 | -26.3388
16.7928 | -34.9311 | -26.3465
16.7989 | -34.9316 | -26.3539
16.805 | -34.9321 | -26.3615
16.8111 | -34.9326 | -26.3694
16.8172 | -34.933 | -26.3771
16.8233 | -34.9335 | -26.3849
16.8294 | -34.934 | -26.3926
16.8356 | -34.9345 | -26.4003
16.8417 | -34.935 | -26.4078
16.8478 | -34.9355 | -26.4155
16.8539 | -34.936 | -26.4231
16.86 | -34.9365 | -26.4307
16.8661 | -34.937 | -26.4381
16.8722 | -34.9375 | -26.446
16.8783 | -34.9379 | -26.4538
16.8844 | -34.9384 | -26.4611
16.8905 | -34.9389 | -26.4686
16.8966 | -34.9394 | -26.4761
16.9027 | -34.9399 | -26.4836
16.9088 | -34.9404 | -26.4913
16.9149 | -34.9408 | -26.4989
16.921 | -34.9414 | -26.5066
16.9271 | -34.9418 | -26.5142
16.9332 | -34.9423 | -26.5218
16.9393 | -34.9428 | -26.5296
16.9454 | -34.9432 | -26.5369
16.9515 | -34.9438 | -26.5444
16.9576 | -34.9443 | -26.5516
16.9637 | -34.9448 | -26.5592
16.9698 | -34.9452 | -26.5668
16.9759 | -34.9457 | -26.5742
16.9821 | -34.9462 | -26.582
16.9882 | -34.9467 | -26.5892
16.9943 | -34.9472 | -26.5967
17.0004 | -34.9477 | -26.6041
17.0065 | -34.9482 | -26.6117
17.0126 | -34.9486 | -26.6192
17.0187 | -34.9491 | -26.6269
17.0248 | -34.9496 | -26.6342
17.0309 | -34.9501 | -26.6419
17.037 | -34.9506 | -26.6492
17.0431 | -34.951 | -26.6565
17.0492 | -34.9515 | -26.6641
17.0553 | -34.952 | -26.6716
17.0614 | -34.9525 | -26.6791
17.0675 | -34.953 | -26.6866
17.0736 | -34.9534 | -26.6943
17.0797 | -34.9539 | -26.7017
17.0858 | -34.9544 | -26.7094
17.0919 | -34.9549 | -26.7169
17.098 | -34.9554 | -26.7242
17.1041 | -34.9559 | -26.7318
17.1102 | -34.9564 | -26.7392
17.1163 | -34.9569 | -26.7468
17.1225 | -34.9574 | -26.7545
17.1286 | -34.9579 | -26.7618
17.1347 | -34.9584 | -26.7694
17.1408 | -34.9588 | -26.7771
17.1469 | -34.9593 | -26.7846
17.153 | -34.9598 | -26.7921
17.1591 | -34.9603 | -26.7995
17.1652 | -34.9608 | -26.8071
17.1713 | -34.9613 | -26.8145
17.1774 | -34.9617 | -26.8222
17.1835 | -34.9622 | -26.8293
17.1896 | -34.9627 | -26.8368
17.1957 | -34.9632 | -26.8443
17.2018 | -34.9637 | -26.8516
17.2079 | -34.9642 | -26.8593
17.214 | -34.9646 | -26.8669
17.2201 | -34.9651 | -26.8744
17.2262 | -34.9656 | -26.8819
17.2323 | -34.966 | -26.8894
17.2384 | -34.9666 | -26.897
17.2445 | -34.9671 | -26.9044
17.2506 | -34.9675 | -26.9119
17.2567 | -34.968 | -26.9197
17.2628 | -34.9685 | -26.9273
17.269 | -34.969 | -26.9349
17.2751 | -34.9695 | -26.9423
17.2812 | -34.97 | -26.9499
17.2873 | -34.9704 | -26.9573
17.2934 | -34.9709 | -26.9648
17.2995 | -34.9714 | -26.9723
17.3056 | -34.9719 | -26.9796
17.3117 | -34.9724 | -26.9871
17.3178 | -34.9729 | -26.9946
17.3239 | -34.9734 | -27.0021
17.33 | -34.9738 | -27.0096
17.3361 | -34.9743 | -27.0172
17.3422 | -34.9748 | -27.0245
17.3483 | -34.9753 | -27.0324
17.3544 | -34.9758 | -27.0399
17.3605 | -34.9762 | -27.0474
17.3666 | -34.9767 | -27.0548
17.3727 | -34.9772 | -27.0625
17.3788 | -34.9777 | -27.0697
17.3849 | -34.9781 | -27.0772
17.391 | -34.9786 | -27.0848
17.3971 | -34.979 | -27.0922
17.4032 | -34.9796 | -27.0997
17.4094 | -34.98 | -27.107
17.4155 | -34.9805 | -27.1143
17.4216 | -34.981 | -27.1216
17.4277 | -34.9815 | -27.1292
17.4338 | -34.982 | -27.1366
17.4399 | -34.9824 | -27.1444
17.446 | -34.9829 | -27.1517
17.4521 | -34.9834 | -27.1591
17.4582 | -34.9839 | -27.1666
17.4643 | -34.9843 | -27.1742
17.4704 | -34.9848 | -27.1817
17.4765 | -34.9853 | -27.1895
17.4826 | -34.9858 | -27.1969
17.4887 | -34.9863 | -27.2044
17.4948 | -34.9868 | -27.2121
17.5009 | -34.9873 | -27.2194
17.507 | -34.9878 | -27.2272
17.5131 | -34.9883 | -27.2347
17.5192 | -34.9887 | -27.242
17.5253 | -34.9892 | -27.2497
17.5314 | -34.9896 | -27.2574
17.5375 | -34.9901 | -27.265
17.5436 | -34.9906 | -27.2725
17.5497 | -34.9911 | -27.2801
17.5559 | -34.9915 | -27.2872
17.562 | -34.992 | -27.2946
17.5681 | -34.9925 | -27.3021
17.5742 | -34.993 | -27.3097
17.5803 | -34.9935 | -27.317
17.5864 | -34.994 | -27.3246
17.5925 | -34.9945 | -27.3321
17.5986 | -34.9949 | -27.3397
17.6047 | -34.9954 | -27.3471
17.6108 | -34.9959 | -27.3549
17.6169 | -34.9964 | -27.3625
17.623 | -34.9969 | -27.37
17.6291 | -34.9974 | -27.3774
17.6352 | -34.9978 | -27.3848
17.6413 | -34.9983 | -27.3922
17.6474 | -34.9988 | -27.3994
17.6535 | -34.9993 | -27.4071
17.6596 | -34.9997 | -27.4141
17.6657 | -35.0002 | -27.4216
17.6718 | -35.0007 | -27.4294
17.6779 | -35.0012 | -27.4367
17.684 | -35.0016 | -27.4442
17.6901 | -35.0021 | -27.4518
17.6963 | -35.0026 | -27.4591
17.7024 | -35.0031 | -27.4664
17.7085 | -35.0036 | -27.4737
17.7146 | -35.0041 | -27.4809
17.7207 | -35.0045 | -27.4882
17.7268 | -35.005 | -27.4956
17.7329 | -35.0055 | -27.5032
17.739 | -35.006 | -27.5107
17.7451 | -35.0064 | -27.5181
17.7512 | -35.0069 | -27.5256
17.7573 | -35.0074 | -27.533
17.7634 | -35.0079 | -27.5401
17.7695 | -35.0084 | -27.5473
17.7756 | -35.0089 | -27.5547
17.7817 | -35.0094 | -27.5621
17.7878 | -35.0099 | -27.5695
17.7939 | -35.0103 | -27.5771
17.8 | -35.0108 | -27.5846
17.8061 | -35.0113 | -27.5924
17.8122 | -35.0118 | -27.5997
17.8183 | -35.0123 | -27.6069
17.8244 | -35.0128 | -27.6144
17.8305 | -35.0132 | -27.6219
17.8366 | -35.0137 | -27.6292
17.8428 | -35.0142 | -27.6367
17.8489 | -35.0147 | -27.6441
17.855 | -35.0151 | -27.6515
17.8611 | -35.0156 | -27.659
17.8672 | -35.0161 | -27.6664
17.8733 | -35.0165 | -27.6739
17.8794 | -35.017 | -27.6811
17.8855 | -35.0175 | -27.6886
17.8916 | -35.018 | -27.6963
17.8977 | -35.0184 | -27.7039
17.9038 | -35.0189 | -27.711
17.9099 | -35.0194 | -27.7184
17.916 | -35.0198 | -27.7258
17.9221 | -35.0203 | -27.7332
17.9282 | -35.0208 | -27.7405
17.9343 | -35.0213 | -27.7481
17.9404 | -35.0217 | -27.7553
17.9465 | -35.0222 | -27.7626
17.9526 | -35.0227 | -27.77
17.9587 | -35.0232 | -27.7774
17.9648 | -35.0237 | -27.7849
17.9709 | -35.0242 | -27.7922
17.977 | -35.0247 | -27.7997
17.9832 | -35.0251 | -27.807
17.9893 | -35.0256 | -27.8145
17.9954 | -35.0261 | -27.822
18.0015 | -35.0265 | -27.8295
18.0076 | -35.0271 | -27.8367
18.0137 | -35.0275 | -27.8442
18.0198 | -35.028 | -27.8516
18.0259 | -35.0285 | -27.8591
18.032 | -35.0289 | -27.8668
18.0381 | -35.0294 | -27.8741
18.0442 | -35.0299 | -27.8815
18.0503 | -35.0303 | -27.889
18.0564 | -35.0308 | -27.8965
18.0625 | -35.0313 | -27.9039
18.0686 | -35.0318 | -27.9112
18.0747 | -35.0322 | -27.919
18.0808 | -35.0326 | -27.9264
18.0869 | -35.0331 | -27.934
18.093 | -35.0336 | -27.9413
18.0991 | -35.0341 | -27.9486
18.1052 | -35.0345 | -27.9559
18.1113 | -35.035 | -27.9631
18.1174 | -35.0355 | -27.9703
18.1236 | -35.0359 | -27.9777
18.1297 | -35.0364 | -27.9851
18.1358 | -35.0368 | -27.9926
18.1419 | -35.0373 | -27.9999
18.148 | -35.0378 | -28.0073
18.1541 | -35.0382 | -28.0145
18.1602 | -35.0387 | -28.0221
18.1663 | -35.0391 | -28.0297
18.1724 | -35.0396 | -28.037
18.1785 | -35.0401 | -28.0443
18.1846 | -35.0405 | -28.0517
18.1907 | -35.041 | -28.0591
18.1968 | -35.0414 | -28.0665
18.2029 | -35.0419 | -28.074
18.209 | -35.0424 | -28.0816
18.2151 | -35.0429 | -28.0891
18.2212 | -35.0433 | -28.0964
18.2273 | -35.0438 | -28.1037
18.2334 | -35.0443 | -28.1109
18.2395 | -35.0447 | -28.1182
18.2456 | -35.0452 | -28.1256
18.2517 | -35.0457 | -28.133
18.2578 | -35.0461 | -28.1401
18.2639 | -35.0466 | -28.1477
18.2701 | -35.0471 | -28.1551
18.2762 | -35.0476 | -28.1622
18.2823 | -35.048 | -28.1697
18.2884 | -35.0485 | -28.1771
18.2945 | -35.049 | -28.1842
18.3006 | -35.0494 | -28.1916
18.3067 | -35.0499 | -28.1989
18.3128 | -35.0504 | -28.2064
18.3189 | -35.0509 | -28.2137
18.325 | -35.0514 | -28.221
18.3311 | -35.0518 | -28.2282
18.3372 | -35.0523 | -28.2356
18.3433 | -35.0528 | -28.2428
18.3494 | -35.0533 | -28.2501
18.3555 | -35.0538 | -28.2575
18.3616 | -35.0542 | -28.2645
18.3677 | -35.0547 | -28.272
18.3738 | -35.0551 | -28.2794
18.3799 | -35.0556 | -28.2867
18.386 | -35.0561 | -28.294
18.3921 | -35.0565 | -28.3014
18.3982 | -35.057 | -28.3085
18.4043 | -35.0574 | -28.316
18.4105 | -35.0579 | -28.3232
18.4166 | -35.0584 | -28.3307
18.4227 | -35.0588 | -28.338
18.4288 | -35.0593 | -28.3452
18.4349 | -35.0597 | -28.3528
18.441 | -35.0602 | -28.3599
18.4471 | -35.0607 | -28.3672
18.4532 | -35.0611 | -28.3746
18.4593 | -35.0616 | -28.382
18.4654 | -35.0621 | -28.389
18.4715 | -35.0625 | -28.3965
18.4776 | -35.0629 | -28.4038
18.4837 | -35.0634 | -28.4113
18.4898 | -35.0639 | -28.4187
18.4959 | -35.0644 | -28.4263
18.502 | -35.0648 | -28.4335
18.5081 | -35.0653 | -28.441
18.5142 | -35.0657 | -28.4484
18.5203 | -35.0662 | -28.4557
18.5264 | -35.0666 | -28.4632
18.5325 | -35.0671 | -28.4706
18.5386 | -35.0676 | -28.478
18.5447 | -35.0681 | -28.4853
18.5508 | -35.0685 | -28.4925
18.557 | -35.069 | -28.4999
18.5631 | -35.0694 | -28.5072
18.5692 | -35.0699 | -28.5147
18.5753 | -35.0704 | -28.522
18.5814 | -35.0708 | -28.5294
18.5875 | -35.0714 | -28.5368
18.5936 | -35.0719 | -28.544
18.5997 | -35.0723 | -28.5515
18.6058 | -35.0728 | -28.5585
18.6119 | -35.0733 | -28.5659
18.618 | -35.0737 | -28.5731
18.6241 | -35.0742 | -28.5808
18.6302 | -35.0747 | -28.5884
18.6363 | -35.0751 | -28.5959
18.6424 | -35.0756 | -28.6032
18.6485 | -35.0761 | -28.6104
18.6546 | -35.0765 | -28.6178
18.6607 | -35.077 | -28.6251
18.6668 | -35.0774 | -28.6322
18.6729 | -35.0779 | -28.6396
18.679 | -35.0784 | -28.647
18.6851 | -35.0789 | -28.6544
18.6912 | -35.0793 | -28.6616
18.6974 | -35.0798 | -28.6689
18.7035 | -35.0803 | -28.6761
18.7096 | -35.0808 | -28.6833
18.7157 | -35.0812 | -28.6907
18.7218 | -35.0817 | -28.6981
18.7279 | -35.0822 | -28.7054
18.734 | -35.0826 | -28.7126
18.7401 | -35.0831 | -28.7198
18.7462 | -35.0835 | -28.7267
18.7523 | -35.084 | -28.734
18.7584 | -35.0845 | -28.7414
18.7645 | -35.085 | -28.7487
18.7706 | -35.0854 | -28.7559
18.7767 | -35.0859 | -28.763
18.7828 | -35.0864 | -28.7699
18.7889 | -35.0868 | -28.7774
18.795 | -35.0873 | -28.7848
18.8011 | -35.0877 | -28.7921
18.8072 | -35.0881 | -28.7994
18.8133 | -35.0886 | -28.8066
18.8194 | -35.0891 | -28.8139
18.8255 | -35.0895 | -28.8212
18.8316 | -35.09 | -28.8285
18.8377 | -35.0905 | -28.8357
18.8439 | -35.0909 | -28.843
18.85 | -35.0914 | -28.8502
18.8561 | -35.0919 | -28.8573
18.8622 | -35.0923 | -28.8647
18.8683 | -35.0928 | -28.8718
18.8744 | -35.0933 | -28.8789
18.8805 | -35.0937 | -28.8862
18.8866 | -35.0942 | -28.8935
18.8927 | -35.0946 | -28.9005
18.8988 | -35.0951 | -28.9078
18.9049 | -35.0956 | -28.9151
18.911 | -35.0961 | -28.9223
18.9171 | -35.0965 | -28.9297
18.9232 | -35.0969 | -28.9371
18.9293 | -35.0974 | -28.9444
18.9354 | -35.0979 | -28.9515
18.9415 | -35.0984 | -28.9588
18.9476 | -35.0988 | -28.9659
18.9537 | -35.0993 | -28.9732
18.9598 | -35.0997 | -28.9803
18.9659 | -35.1002 | -28.9875
18.972 | -35.1006 | -28.9948
18.9781 | -35.1011 | -29.0019
18.9843 | -35.1015 | -29.009
18.9904 | -35.102 | -29.0161
18.9965 | -35.1025 | -29.0232
19.0026 | -35.1029 | -29.0304
19.0087 | -35.1034 | -29.0376
19.0148 | -35.1039 | -29.0449
19.0209 | -35.1043 | -29.0523
19.027 | -35.1048 | -29.0595
19.0331 | -35.1052 | -29.0666
19.0392 | -35.1057 | -29.0742
19.0453 | -35.1062 | -29.0814
19.0514 | -35.1066 | -29.0885
19.0575 | -35.1071 | -29.0955
19.0636 | -35.1075 | -29.1025
19.0697 | -35.108 | -29.1096
19.0758 | -35.1084 | -29.117
19.0819 | -35.1089 | -29.1241
19.088 | -35.1093 | -29.1312
19.0941 | -35.1098 | -29.1385
19.1002 | -35.1103 | -29.1455
19.1063 | -35.1107 | -29.1524
19.1124 | -35.1111 | -29.1596
19.1185 | -35.1116 | -29.1667
19.1246 | -35.1121 | -29.1738
19.1308 | -35.1126 | -29.1812
19.1369 | -35.113 | -29.1885
19.143 | -35.1135 | -29.1959
19.1491 | -35.114 | -29.2032
19.1552 | -35.1144 | -29.2103
19.1613 | -35.1148 | -29.2175
19.1674 | -35.1153 | -29.2249
19.1735 | -35.1158 | -29.232
19.1796 | -35.1162 | -29.2392
19.1857 | -35.1167 | -29.2462
19.1918 | -35.1172 | -29.2533
19.1979 | -35.1176 | -29.2604
19.204 | -35.118 | -29.2678
19.2101 | -35.1185 | -29.2748
19.2162 | -35.119 | -29.282
19.2223 | -35.1194 | -29.2893
19.2284 | -35.1199 | -29.2967
19.2345 | -35.1203 | -29.304
19.2406 | -35.1208 | -29.3113
19.2467 | -35.1212 | -29.3186
19.2528 | -35.1217 | -29.3261
19.2589 | -35.1221 | -29.333
19.265 | -35.1226 | -29.34
19.2712 | -35.123 | -29.3471
19.2773 | -35.1235 | -29.3545
19.2834 | -35.124 | -29.3616
19.2895 | -35.1245 | -29.3688
19.2956 | -35.1249 | -29.3761
19.3017 | -35.1254 | -29.3832
19.3078 | -35.1258 | -29.3904
19.3139 | -35.1263 | -29.3974
19.32 | -35.1268 | -29.4045
19.3261 | -35.1272 | -29.4115
19.3322 | -35.1277 | -29.4186
19.3383 | -35.1282 | -29.4257
19.3444 | -35.1286 | -29.4327
19.3505 | -35.129 | -29.4402
19.3566 | -35.1295 | -29.4475
19.3627 | -35.13 | -29.4543
19.3688 | -35.1305 | -29.4613
19.3749 | -35.1309 | -29.4687
19.381 | -35.1314 | -29.476
19.3871 | -35.1319 | -29.4832
19.3932 | -35.1323 | -29.4903
19.3993 | -35.1328 | -29.4975
19.4054 | -35.1332 | -29.5047
19.4115 | -35.1337 | -29.5119
19.4177 | -35.1341 | -29.5187
19.4238 | -35.1346 | -29.5257
19.4299 | -35.135 | -29.5327
19.436 | -35.1355 | -29.5402
19.4421 | -35.136 | -29.5471
19.4482 | -35.1364 | -29.5541
19.4543 | -35.1369 | -29.5608
19.4604 | -35.1373 | -29.5679
19.4665 | -35.1378 | -29.5748
19.4726 | -35.1382 | -29.5818
19.4787 | -35.1387 | -29.5889
19.4848 | -35.1392 | -29.596
19.4909 | -35.1396 | -29.6027
19.497 | -35.1401 | -29.6098
19.5031 | -35.1405 | -29.6167
19.5092 | -35.1409 | -29.6235
19.5153 | -35.1415 | -29.6306
19.5214 | -35.1419 | -29.6378
19.5275 | -35.1424 | -29.6451
19.5336 | -35.1428 | -29.6523
19.5397 | -35.1432 | -29.6594
19.5458 | -35.1437 | -29.6662
19.5519 | -35.1441 | -29.6735
19.5581 | -35.1446 | -29.6804
19.5642 | -35.145 | -29.6874
19.5703 | -35.1455 | -29.6946
19.5764 | -35.1459 | -29.7018
19.5825 | -35.1464 | -29.7087
19.5886 | -35.1469 | -29.7158
19.5947 | -35.1473 | -29.7226
19.6008 | -35.1478 | -29.7296
19.6069 | -35.1483 | -29.7366
19.613 | -35.1488 | -29.744
19.6191 | -35.1492 | -29.7506
19.6252 | -35.1497 | -29.7579
19.6313 | -35.1502 | -29.7646
19.6374 | -35.1506 | -29.7717
19.6435 | -35.1511 | -29.7787
19.6496 | -35.1515 | -29.7858
19.6557 | -35.152 | -29.7929
19.6618 | -35.1524 | -29.7999
19.6679 | -35.1528 | -29.8065
19.674 | -35.1533 | -29.8135
19.6801 | -35.1537 | -29.8204
19.6862 | -35.1542 | -29.8275
19.6923 | -35.1546 | -29.8345
19.6984 | -35.1551 | -29.8409
19.7046 | -35.1556 | -29.8477
19.7107 | -35.1561 | -29.8542
19.7168 | -35.1565 | -29.8612
19.7229 | -35.157 | -29.8681
19.729 | -35.1574 | -29.875
19.7351 | -35.1579 | -29.8817
19.7412 | -35.1584 | -29.8886
19.7473 | -35.1588 | -29.8956
19.7534 | -35.1593 | -29.9026
19.7595 | -35.1597 | -29.9098
19.7656 | -35.1602 | -29.9168
19.7717 | -35.1606 | -29.9237
19.7778 | -35.1611 | -29.9308
19.7839 | -35.1616 | -29.9377
19.79 | -35.1621 | -29.9448
19.7961 | -35.1624 | -29.952
19.8022 | -35.1629 | -29.9582
19.8083 | -35.1634 | -29.965
19.8144 | -35.1638 | -29.9716
19.8205 | -35.1641 | -29.9785
19.8266 | -35.1645 | -29.985
19.8327 | -35.165 | -29.992
19.8388 | -35.1655 | -29.9983
19.845 | -35.1659 | -30.0051
19.8511 | -35.1664 | -30.0121
19.8572 | -35.1668 | -30.019
19.8633 | -35.1673 | -30.0262
19.8694 | -35.1677 | -30.033
19.8755 | -35.1682 | -30.0398
19.8816 | -35.1686 | -30.0467
19.8877 | -35.1691 | -30.0541
19.8938 | -35.1695 | -30.0614
19.8999 | -35.17 | -30.0684
19.906 | -35.1704 | -30.0755
19.9121 | -35.1709 | -30.0823
19.9182 | -35.1714 | -30.0889
19.9243 | -35.1718 | -30.0958
19.9304 | -35.1723 | -30.1028
19.9365 | -35.1728 | -30.1097
19.9426 | -35.1733 | -30.1163
19.9487 | -35.1737 | -30.1229
19.9548 | -35.1741 | -30.1298
19.9609 | -35.1746 | -30.1366
19.967 | -35.1751 | -30.1435
19.9731 | -35.1756 | -30.15
19.9792 | -35.1761 | -30.1562
19.9853 | -35.1765 | -30.1631
19.9915 | -35.1771 | -30.17
19.9976 | -35.1775 | -30.1766
20.0037 | -35.1779 | -30.1837
20.0098 | -35.1783 | -30.1903
20.0159 | -35.1788 | -30.1971
20.022 | -35.1792 | -30.2039
20.0281 | -35.1797 | -30.2108
20.0342 | -35.1802 | -30.2176
20.0403 | -35.1807 | -30.2242
20.0464 | -35.1811 | -30.2309
20.0525 | -35.1816 | -30.2373
20.0586 | -35.182 | -30.2443
20.0647 | -35.1824 | -30.2509
20.0708 | -35.1829 | -30.258
20.0769 | -35.1835 | -30.2652
20.083 | -35.184 | -30.2716
20.0891 | -35.1844 | -30.2786
20.0952 | -35.1849 | -30.2853
20.1013 | -35.1853 | -30.2922
20.1074 | -35.1857 | -30.2986
20.1135 | -35.1862 | -30.3055
20.1196 | -35.1866 | -30.3121
20.1257 | -35.1871 | -30.3191
20.1319 | -35.1875 | -30.3258
20.138 | -35.1879 | -30.333
20.1441 | -35.1884 | -30.34
20.1502 | -35.1888 | -30.347
20.1563 | -35.1893 | -30.3536
20.1624 | -35.1899 | -30.3604
20.1685 | -35.1903 | -30.367
20.1746 | -35.1908 | -30.3735
20.1807 | -35.1913 | -30.3802
20.1868 | -35.1918 | -30.3871
20.1929 | -35.1923 | -30.3937
20.199 | -35.1927 | -30.4007
20.2051 | -35.1932 | -30.4073
20.2112 | -35.1936 | -30.4141
20.2173 | -35.1941 | -30.4207
20.2234 | -35.1946 | -30.4271
20.2295 | -35.1951 | -30.4338
20.2356 | -35.1956 | -30.4406
20.2417 | -35.196 | -30.4473
20.2478 | -35.1965 | -30.4537
20.2539 | -35.1968 | -30.4611
20.26 | -35.1972 | -30.4671
20.2661 | -35.1977 | -30.4741
20.2723 | -35.1982 | -30.4807
20.2784 | -35.1987 | -30.4873
20.2845 | -35.1992 | -30.4935
20.2906 | -35.1997 | -30.5005
20.2967 | -35.2002 | -30.506
20.3028 | -35.2006 | -30.5126
20.3089 | -35.2011 | -30.5195
20.315 | -35.2015 | -30.5264
20.3211 | -35.202 | -30.5332
20.3272 | -35.2025 | -30.5395
20.3333 | -35.2029 | -30.5462
20.3394 | -35.2034 | -30.5531
20.3455 | -35.2038 | -30.5598
20.3516 | -35.2042 | -30.5667
20.3577 | -35.2047 | -30.5736
20.3638 | -35.2051 | -30.5803
20.3699 | -35.2056 | -30.5864
20.376 | -35.206 | -30.5925
20.3821 | -35.2065 | -30.5994
20.3882 | -35.207 | -30.6057
20.3943 | -35.2074 | -30.6122
20.4004 | -35.2078 | -30.6186
20.4065 | -35.2083 | -30.6254
20.4126 | -35.2087 | -30.6321
20.4188 | -35.2092 | -30.6386
20.4249 | -35.2097 | -30.6457
20.431 | -35.2101 | -30.6526
20.4371 | -35.2107 | -30.6591
20.4432 | -35.2111 | -30.6652
20.4493 | -35.2115 | -30.6723
20.4554 | -35.2119 | -30.6787
20.4615 | -35.2124 | -30.6857
20.4676 | -35.2128 | -30.6926
20.4737 | -35.2132 | -30.6991
20.4798 | -35.2137 | -30.7053
20.4859 | -35.2141 | -30.7121
20.492 | -35.2146 | -30.7187
20.4981 | -35.2151 | -30.7252
20.5042 | -35.2156 | -30.7322
20.5103 | -35.216 | -30.7389
20.5164 | -35.2164 | -30.7453
20.5225 | -35.2169 | -30.7515
20.5286 | -35.2175 | -30.7582
20.5347 | -35.2179 | -30.7645
20.5408 | -35.2183 | -30.7707
20.5469 | -35.2187 | -30.7774
20.553 | -35.2192 | -30.7839
20.5592 | -35.2196 | -30.7906
20.5653 | -35.2201 | -30.797
20.5714 | -35.2205 | -30.8034
20.5775 | -35.221 | -30.8098
20.5836 | -35.2215 | -30.8163
20.5897 | -35.222 | -30.8232
20.5958 | -35.2224 | -30.8288
20.6019 | -35.2228 | -30.8352
20.608 | -35.2231 | -30.8413
20.6141 | -35.2236 | -30.8488
20.6202 | -35.2242 | -30.8549
20.6263 | -35.2246 | -30.8611
20.6324 | -35.2251 | -30.8677
20.6385 | -35.2255 | -30.8744
20.6446 | -35.226 | -30.8812
20.6507 | -35.2264 | -30.8877
20.6568 | -35.2268 | -30.8943
20.6629 | -35.2273 | -30.9008
20.669 | -35.2277 | -30.9072
20.6751 | -35.2281 | -30.9139
20.6812 | -35.2285 | -30.92
20.6873 | -35.2289 | -30.9266
20.6934 | -35.2293 | -30.9333
20.6995 | -35.2298 | -30.9398
20.7057 | -35.2304 | -30.9464
20.7118 | -35.2307 | -30.9525
20.7179 | -35.2312 | -30.9588
20.724 | -35.2317 | -30.9654
20.7301 | -35.2321 | -30.9712
20.7362 | -35.2327 | -30.9774
20.7423 | -35.2332 | -30.9839
20.7484 | -35.2336 | -30.9909
20.7545 | -35.2341 | -30.9954
20.7606 | -35.2345 | -31.0017
20.7667 | -35.235 | -31.0082
20.7728 | -35.2355 | -31.0149
20.7789 | -35.236 | -31.0213
20.785 | -35.2364 | -31.0282
20.7911 | -35.2368 | -31.0347
20.7972 | -35.2371 | -31.0403
20.8033 | -35.2376 | -31.0469
20.8094 | -35.238 | -31.0531
20.8155 | -35.2385 | -31.0595
20.8216 | -35.2389 | -31.0664
20.8277 | -35.2394 | -31.0725
20.8338 | -35.2398 | -31.0789
20.8399 | -35.2403 | -31.085
20.8461 | -35.2407 | -31.0915
20.8522 | -35.2412 | -31.0978
20.8583 | -35.2417 | -31.1037
20.8644 | -35.2422 | -31.1101
20.8705 | -35.2426 | -31.1164
20.8766 | -35.243 | -31.1231
20.8827 | -35.2435 | -31.1294
20.8888 | -35.244 | -31.1365
20.8949 | -35.2445 | -31.1421
20.901 | -35.2449 | -31.1478
20.9071 | -35.2455 | -31.1545
20.9132 | -35.2461 | -31.1614
20.9193 | -35.2465 | -31.1682
20.9254 | -35.247 | -31.1746
20.9315 | -35.2475 | -31.1805
20.9376 | -35.248 | -31.1865
20.9437 | -35.2484 | -31.1929
20.9498 | -35.2489 | -31.1992
20.9559 | -35.2493 | -31.2054
20.962 | -35.2498 | -31.2114
20.9681 | -35.2502 | -31.2175
20.9742 | -35.2507 | -31.2235
20.9803 | -35.2513 | -31.23
20.9864 | -35.2518 | -31.2357
20.9926 | -35.2523 | -31.2422
20.9987 | -35.2528 | -31.2489
21.0048 | -35.2533 | -31.2555
21.0109 | -35.2537 | -31.2618
21.017 | -35.2542 | -31.2684
21.0231 | -35.2547 | -31.2742
21.0292 | -35.2551 | -31.2804
21.0353 | -35.2556 | -31.2862
21.0414 | -35.2559 | -31.2923
21.0475 | -35.2564 | -31.2988
21.0536 | -35.2569 | -31.305
21.0597 | -35.2573 | -31.3105
21.0658 | -35.2579 | -31.3179
21.0719 | -35.2584 | -31.3232
21.078 | -35.2589 | -31.33
21.0841 | -35.2592 | -31.3367
21.0902 | -35.2597 | -31.343
21.0963 | -35.2601 | -31.3493
21.1024 | -35.2607 | -31.356
21.1085 | -35.2611 | -31.3628
21.1146 | -35.2615 | -31.3689
21.1207 | -35.262 | -31.3754
21.1268 | -35.2623 | -31.3819
21.133 | -35.2627 | -31.3886
21.1391 | -35.2631 | -31.3952
21.1452 | -35.2636 | -31.4011
21.1513 | -35.2641 | -31.4073
21.1574 | -35.2645 | -31.414
21.1635 | -35.265 | -31.4213
21.1696 | -35.2655 | -31.4279
21.1757 | -35.2659 | -31.434
21.1818 | -35.2665 | -31.4399
21.1879 | -35.2669 | -31.4464
21.194 | -35.2675 | -31.4527
21.2001 | -35.268 | -31.4584
21.2062 | -35.2685 | -31.4642
21.2123 | -35.2689 | -31.4695
21.2184 | -35.2694 | -31.4744
21.2245 | -35.2698 | -31.4806
21.2306 | -35.2702 | -31.4868
21.2367 | -35.2706 | -31.493
21.2428 | -35.2711 | -31.4987
21.2489 | -35.2715 | -31.505
21.255 | -35.272 | -31.5116
21.2611 | -35.2724 | -31.5177
21.2672 | -35.2728 | -31.5238
21.2733 | -35.2733 | -31.5296
21.2795 | -35.2738 | -31.5356
21.2856 | -35.2742 | -31.5417
21.2917 | -35.2747 | -31.5475
21.2978 | -35.2751 | -31.5534
21.3039 | -35.2755 | -31.559
21.31 | -35.276 | -31.5647
21.3161 | -35.2765 | -31.5712
21.3222 | -35.2769 | -31.5775
21.3283 | -35.2774 | -31.5835
21.3344 | -35.2778 | -31.5901
21.3405 | -35.2783 | -31.5969
21.3466 | -35.2787 | -31.6035
21.3527 | -35.2792 | -31.6091
21.3588 | -35.2797 | -31.6151
21.3649 | -35.2801 | -31.6214
21.371 | -35.2805 | -31.6284
21.3771 | -35.281 | -31.635
21.3832 | -35.2814 | -31.641
21.3893 | -35.2818 | -31.6468
21.3954 | -35.2822 | -31.6533
21.4015 | -35.2827 | -31.6595
21.4076 | -35.2832 | -31.6656
21.4137 | -35.2837 | -31.672
21.4199 | -35.2841 | -31.678
21.426 | -35.2845 | -31.6842
21.4321 | -35.2848 | -31.6912
21.4382 | -35.2854 | -31.697
21.4443 | -35.2858 | -31.703
21.4504 | -35.2863 | -31.7095
21.4565 | -35.2867 | -31.715
21.4626 | -35.2872 | -31.7209
21.4687 | -35.2876 | -31.7272
21.4748 | -35.2881 | -31.7331
21.4809 | -35.2885 | -31.7391
21.487 | -35.289 | -31.7452
21.4931 | -35.2894 | -31.751
21.4992 | -35.2899 | -31.7569
21.5053 | -35.2903 | -31.7629
21.5114 | -35.2908 | -31.7691
21.5175 | -35.2912 | -31.7749
21.5236 | -35.2915 | -31.7811
21.5297 | -35.2919 | -31.7869
21.5358 | -35.2924 | -31.7921
21.5419 | -35.2929 | -31.798
21.548 | -35.2934 | -31.8041
21.5541 | -35.2938 | -31.8106
21.5602 | -35.2943 | -31.8167
21.5664 | -35.2947 | -31.8223
21.5725 | -35.2951 | -31.8282
21.5786 | -35.2956 | -31.8346
21.5847 | -35.296 | -31.8407
21.5908 | -35.2965 | -31.8467
21.5969 | -35.2968 | -31.8523
21.603 | -35.2972 | -31.8588
21.6091 | -35.2976 | -31.865
21.6152 | -35.298 | -31.8719
21.6213 | -35.2985 | -31.879
21.6274 | -35.2989 | -31.8853
21.6335 | -35.2993 | -31.8905
21.6396 | -35.2997 | -31.8964
21.6457 | -35.3001 | -31.9027
21.6518 | -35.3006 | -31.9086
21.6579 | -35.301 | -31.9146
21.664 | -35.3014 | -31.921
21.6701 | -35.3019 | -31.9264
21.6762 | -35.3023 | -31.9331
21.6823 | -35.3027 | -31.9405
21.6884 | -35.3033 | -31.9469
21.6945 | -35.3038 | -31.9529
21.7006 | -35.3043 | -31.9585
21.7068 | -35.3049 | -31.964
21.7129 | -35.3054 | -31.9703
21.719 | -35.3058 | -31.9762
21.7251 | -35.3063 | -31.9823
21.7312 | -35.3067 | -31.9881
21.7373 | -35.3071 | -31.9939
21.7434 | -35.3076 | -31.9996
21.7495 | -35.3081 | -32.0057
21.7556 | -35.3085 | -32.0119
21.7617 | -35.3089 | -32.0181
21.7678 | -35.3093 | -32.024
21.7739 | -35.3098 | -32.0296
21.78 | -35.3102 | -32.0353
21.7861 | -35.3106 | -32.0415
21.7922 | -35.3111 | -32.0475
21.7983 | -35.3114 | -32.0529
21.8044 | -35.3119 | -32.0588
21.8105 | -35.3123 | -32.0645
21.8166 | -35.3128 | -32.0708
21.8227 | -35.3132 | -32.0769
21.8288 | -35.3138 | -32.0828
21.8349 | -35.3142 | -32.0883
21.841 | -35.3148 | -32.0941
21.8471 | -35.3154 | -32.1002
21.8533 | -35.3158 | -32.1063
21.8594 | -35.3163 | -32.1121
21.8655 | -35.3168 | -32.1182
21.8716 | -35.3173 | -32.1243
21.8777 | -35.3178 | -32.1304
21.8838 | -35.3181 | -32.1369
21.8899 | -35.3185 | -32.1437
21.896 | -35.319 | -32.1494
21.9021 | -35.3194 | -32.1553
21.9082 | -35.3199 | -32.161
21.9143 | -35.3203 | -32.1663
21.9204 | -35.3207 | -32.1721
21.9265 | -35.3211 | -32.1778
21.9326 | -35.3216 | -32.1835
21.9387 | -35.322 | -32.1893
21.9448 | -35.3225 | -32.1957
21.9509 | -35.3229 | -32.2019
21.957 | -35.3234 | -32.208
21.9631 | -35.3238 | -32.2136
21.9692 | -35.3242 | -32.2196
21.9753 | -35.3247 | -32.2257
21.9814 | -35.3253 | -32.2323
21.9875 | -35.3257 | -32.2382
21.9937 | -35.3263 | -32.2444
21.9998 | -35.3268 | -32.2493
22.0059 | -35.3272 | -32.2552
22.012 | -35.3277 | -32.2609
22.0181 | -35.3281 | -32.2665
22.0242 | -35.3285 | -32.2723
22.0303 | -35.329 | -32.2783
22.0364 | -35.3295 | -32.2845
22.0425 | -35.3299 | -32.2903
22.0486 | -35.3303 | -32.296
22.0547 | -35.3307 | -32.302
22.0608 | -35.3311 | -32.308
22.0669 | -35.3316 | -32.3143
22.073 | -35.3319 | -32.32
22.0791 | -35.3324 | -32.326
22.0852 | -35.3328 | -32.3321
22.0913 | -35.3333 | -32.3377
22.0974 | -35.3337 | -32.3435
22.1035 | -35.3341 | -32.3497
22.1096 | -35.3346 | -32.3557
22.1157 | -35.335 | -32.3615
22.1218 | -35.3355 | -32.3674
22.1279 | -35.336 | -32.3736
22.134 | -35.3364 | -32.3787
22.1402 | -35.3369 | -32.3852
22.1463 | -35.3373 | -32.3902
22.1524 | -35.3379 | -32.3955
22.1585 | -35.3384 | -32.4016
22.1646 | -35.3389 | -32.4068
22.1707 | -35.3395 | -32.413
22.1768 | -35.34 | -32.4188
22.1829 | -35.3405 | -32.4251
22.189 | -35.3409 | -32.4311
22.1951 | -35.3414 | -32.4367
22.2012 | -35.3418 | -32.4428
22.2073 | -35.3423 | -32.4484
22.2134 | -35.3427 | -32.4541
22.2195 | -35.3432 | -32.4599
22.2256 | -35.3436 | -32.4659
22.2317 | -35.3441 | -32.4721
22.2378 | -35.3446 | -32.4776
22.2439 | -35.3449 | -32.4836
22.25 | -35.3455 | -32.4897
22.2561 | -35.3458 | -32.4955
22.2622 | -35.3464 | -32.5015
22.2683 | -35.3469 | -32.5077
22.2744 | -35.3473 | -32.5134
22.2806 | -35.3478 | -32.5191
22.2867 | -35.3483 | -32.5253
22.2928 | -35.3487 | -32.5318
22.2989 | -35.3491 | -32.5371
22.305 | -35.3493 | -32.5439
22.3111 | -35.3499 | -32.5501
22.3172 | -35.3503 | -32.5558
22.3233 | -35.3508 | -32.5621
22.3294 | -35.3512 | -32.5684
22.3355 | -35.3514 | -32.5743
22.3416 | -35.3518 | -32.5808
22.3477 | -35.3522 | -32.587
22.3538 | -35.3526 | -32.5929
22.3599 | -35.353 | -32.5984
22.366 | -35.3535 | -32.6041
22.3721 | -35.354 | -32.6093
22.3782 | -35.3544 | -32.6152
22.3843 | -35.3549 | -32.6212
22.3904 | -35.3553 | -32.6268
22.3965 | -35.3557 | -32.6326
22.4026 | -35.3562 | -32.6387
22.4087 | -35.3566 | -32.6442
22.4148 | -35.357 | -32.6496
22.4209 | -35.3574 | -32.6554
22.4271 | -35.3579 | -32.6609
22.4332 | -35.3583 | -32.6669
22.4393 | -35.3588 | -32.672
22.4454 | -35.3592 | -32.6772
22.4515 | -35.3596 | -32.6841
22.4576 | -35.36 | -32.6896
22.4637 | -35.3605 | -32.6953
22.4698 | -35.3609 | -32.7015
22.4759 | -35.3613 | -32.7065
22.482 | -35.3618 | -32.7125
22.4881 | -35.3621 | -32.7183
22.4942 | -35.3626 | -32.7238
22.5003 | -35.3629 | -32.7293
22.5064 | -35.3634 | -32.7354
22.5125 | -35.3638 | -32.741
22.5186 | -35.3643 | -32.7469
22.5247 | -35.3646 | -32.7525
22.5308 | -35.365 | -32.7582
22.5369 | -35.3655 | -32.7643
22.543 | -35.3659 | -32.7703
22.5491 | -35.3664 | -32.7758
22.5552 | -35.3668 | -32.7816
22.5613 | -35.3673 | -32.7876
22.5675 | -35.3678 | -32.7932
22.5736 | -35.3682 | -32.7989
22.5797 | -35.3686 | -32.8051
22.5858 | -35.369 | -32.8111
22.5919 | -35.3695 | -32.817
22.598 | -35.37 | -32.8229
22.6041 | -35.3704 | -32.8287
22.6102 | -35.3708 | -32.835
22.6163 | -35.3712 | -32.8409
22.6224 | -35.3716 | -32.8468
22.6285 | -35.3721 | -32.8526
22.6346 | -35.3724 | -32.8587
22.6407 | -35.3729 | -32.864
22.6468 | -35.3734 | -32.8699
22.6529 | -35.3738 | -32.8757
22.659 | -35.3743 | -32.8813
22.6651 | -35.3747 | -32.8874
22.6712 | -35.3751 | -32.8932
22.6773 | -35.3756 | -32.8991
22.6834 | -35.3761 | -32.9051
22.6895 | -35.3764 | -32.9108
22.6956 | -35.3768 | -32.9167
22.7017 | -35.3773 | -32.9224
22.7079 | -35.3778 | -32.9285
22.714 | -35.3782 | -32.9344
22.7201 | -35.3786 | -32.9398
22.7262 | -35.3791 | -32.946
22.7323 | -35.3795 | -32.9517
22.7384 | -35.3799 | -32.9576
22.7445 | -35.3803 | -32.9631
22.7506 | -35.3808 | -32.969
22.7567 | -35.3812 | -32.9749
22.7628 | -35.3817 | -32.9808
22.7689 | -35.3821 | -32.9862
22.775 | -35.3826 | -32.9924
22.7811 | -35.3831 | -32.9983
22.7872 | -35.3836 | -33.0043
22.7933 | -35.3841 | -33.0097
22.7994 | -35.3845 | -33.0155
22.8055 | -35.3849 | -33.0216
22.8116 | -35.3854 | -33.0273
22.8177 | -35.3858 | -33.0331
22.8238 | -35.3863 | -33.0391
22.8299 | -35.3868 | -33.0449
22.836 | -35.3872 | -33.0509
22.8421 | -35.3877 | -33.0567
22.8482 | -35.3881 | -33.0624
22.8544 | -35.3885 | -33.0681
22.8605 | -35.3891 | -33.0739
22.8666 | -35.3895 | -33.0802
22.8727 | -35.39 | -33.0859
22.8788 | -35.3904 | -33.0914
22.8849 | -35.3908 | -33.0969
22.891 | -35.3912 | -33.1031
22.8971 | -35.3916 | -33.1088
22.9032 | -35.392 | -33.1143
22.9093 | -35.3925 | -33.1201
22.9154 | -35.393 | -33.1257
22.9215 | -35.3933 | -33.1318
22.9276 | -35.3939 | -33.138
22.9337 | -35.3943 | -33.1439
22.9398 | -35.3948 | -33.1499
22.9459 | -35.3953 | -33.1558
22.952 | -35.3957 | -33.1616
22.9581 | -35.3961 | -33.1675
22.9642 | -35.3965 | -33.1736
22.9703 | -35.397 | -33.1792
22.9764 | -35.3975 | -33.1854
22.9825 | -35.3979 | -33.1911
22.9886 | -35.3983 | -33.197
22.9948 | -35.3987 | -33.2032
23.0009 | -35.3991 | -33.2083
23.007 | -35.3995 | -33.2139
23.0131 | -35.4 | -33.2198
23.0192 | -35.4004 | -33.2254
23.0253 | -35.4008 | -33.231
23.0314 | -35.4013 | -33.2368
23.0375 | -35.4018 | -33.2426
23.0436 | -35.4021 | -33.2482
23.0497 | -35.4026 | -33.254
23.0558 | -35.403 | -33.2595
23.0619 | -35.4034 | -33.2657
23.068 | -35.4039 | -33.2717
23.0741 | -35.4044 | -33.2772
23.0802 | -35.4047 | -33.2828
23.0863 | -35.4052 | -33.2883
23.0924 | -35.4056 | -33.2943
23.0985 | -35.4061 | -33.2999
23.1046 | -35.4065 | -33.3056
23.1107 | -35.4069 | -33.3112
23.1168 | -35.4073 | -33.3168
23.1229 | -35.4077 | -33.3225
23.129 | -35.4081 | -33.3282
23.1351 | -35.4086 | -33.3336
23.1413 | -35.409 | -33.3393
23.1474 | -35.4094 | -33.3449
23.1535 | -35.4098 | -33.3505
23.1596 | -35.4103 | -33.356
23.1657 | -35.4107 | -33.362
23.1718 | -35.4112 | -33.3676
23.1779 | -35.4116 | -33.3731
23.184 | -35.412 | -33.3782
23.1901 | -35.4124 | -33.3843
23.1962 | -35.4129 | -33.3898
23.2023 | -35.4133 | -33.3956
23.2084 | -35.4137 | -33.4012
23.2145 | -35.4141 | -33.4068
23.2206 | -35.4144 | -33.4122
23.2267 | -35.4149 | -33.4182
23.2328 | -35.4152 | -33.424
23.2389 | -35.4156 | -33.4292
23.245 | -35.4161 | -33.4348
23.2511 | -35.4165 | -33.4407
23.2572 | -35.417 | -33.4464
23.2633 | -35.4174 | -33.4521
23.2694 | -35.4177 | -33.4581
23.2755 | -35.4182 | -33.4643
23.2817 | -35.4187 | -33.4701
23.2878 | -35.4192 | -33.4757
23.2939 | -35.4196 | -33.4819
23.3 | -35.4201 | -33.4876
23.3061 | -35.4205 | -33.4932
23.3122 | -35.421 | -33.4985
23.3183 | -35.4214 | -33.5041
23.3244 | -35.4218 | -33.5099
23.3305 | -35.4222 | -33.5158
23.3366 | -35.4226 | -33.5215
23.3427 | -35.4231 | -33.5273
23.3488 | -35.4235 | -33.533
23.3549 | -35.4239 | -33.5389
23.361 | -35.4244 | -33.5446
23.3671 | -35.4248 | -33.5506
23.3732 | -35.4252 | -33.5564
23.3793 | -35.4256 | -33.5621
23.3854 | -35.426 | -33.5683
23.3915 | -35.4265 | -33.5738
23.3976 | -35.427 | -33.5791
23.4037 | -35.4273 | -33.5848
23.4098 | -35.4278 | -33.5907
23.4159 | -35.4282 | -33.5963
23.422 | -35.4287 | -33.6024
23.4282 | -35.4291 | -33.6078
23.4343 | -35.4295 | -33.6134
23.4404 | -35.4299 | -33.6189
23.4465 | -35.4303 | -33.6249
23.4526 | -35.4308 | -33.6307
23.4587 | -35.4312 | -33.6363
23.4648 | -35.4317 | -33.642
23.4709 | -35.432 | -33.6481
23.477 | -35.4325 | -33.6538
23.4831 | -35.4329 | -33.6595
23.4892 | -35.4333 | -33.6653
23.4953 | -35.4338 | -33.6711
23.5014 | -35.4342 | -33.6772
23.5075 | -35.4347 | -33.683
23.5136 | -35.4351 | -33.6886
23.5197 | -35.4355 | -33.6944
23.5258 | -35.436 | -33.7
23.5319 | -35.4364 | -33.7052
23.538 | -35.4368 | -33.711
23.5441 | -35.4372 | -33.7163
23.5502 | -35.4376 | -33.722
23.5563 | -35.4381 | -33.728
23.5624 | -35.4385 | -33.7333
23.5686 | -35.4389 | -33.7389
23.5747 | -35.4394 | -33.7445
23.5808 | -35.4398 | -33.7503
23.5869 | -35.4402 | -33.756
23.593 | -35.4407 | -33.7617
23.5991 | -35.4412 | -33.7676
23.6052 | -35.4415 | -33.7739
23.6113 | -35.4419 | -33.7797
23.6174 | -35.4424 | -33.7856
23.6235 | -35.4428 | -33.7912
23.6296 | -35.4433 | -33.7968
23.6357 | -35.4437 | -33.8029
23.6418 | -35.4441 | -33.8086
23.6479 | -35.4445 | -33.8141
23.654 | -35.445 | -33.8197
23.6601 | -35.4454 | -33.8252
23.6662 | -35.4458 | -33.8309
23.6723 | -35.4461 | -33.8369
23.6784 | -35.4465 | -33.8425
23.6845 | -35.4469 | -33.8486
23.6906 | -35.4473 | -33.8544
23.6967 | -35.4477 | -33.8602
23.7028 | -35.4481 | -33.8662
23.7089 | -35.4484 | -33.8723
23.7151 | -35.4489 | -33.8781
23.7212 | -35.4493 | -33.8837
23.7273 | -35.4497 | -33.8896
23.7334 | -35.4502 | -33.8947
23.7395 | -35.4506 | -33.9003
23.7456 | -35.451 | -33.9062
23.7517 | -35.4514 | -33.9119
23.7578 | -35.4518 | -33.9173
23.7639 | -35.4522 | -33.9233
23.77 | -35.4526 | -33.9288
23.7761 | -35.453 | -33.9346
23.7822 | -35.4534 | -33.9401
23.7883 | -35.4538 | -33.9459
23.7944 | -35.4542 | -33.9518
23.8005 | -35.4546 | -33.9577
23.8066 | -35.4551 | -33.9633
23.8127 | -35.4555 | -33.9695
23.8188 | -35.456 | -33.9753
23.8249 | -35.4564 | -33.9812
23.831 | -35.4568 | -33.9871
23.8371 | -35.4572 | -33.9933
23.8432 | -35.4576 | -33.9987
23.8493 | -35.458 | -34.0041
23.8555 | -35.4584 | -34.01
23.8616 | -35.4587 | -34.0158
23.8677 | -35.4591 | -34.0219
23.8738 | -35.4595 | -34.0278
23.8799 | -35.46 | -34.0336
23.886 | -35.4603 | -34.0391
23.8921 | -35.4608 | -34.0448
23.8982 | -35.4612 | -34.0508
23.9043 | -35.4616 | -34.0562
23.9104 | -35.462 | -34.0621
23.9165 | -35.4625 | -34.0675
23.9226 | -35.4629 | -34.0733
23.9287 | -35.4633 | -34.0795
23.9348 | -35.4637 | -34.0854
23.9409 | -35.4641 | -34.0912
23.947 | -35.4645 | -34.0969
23.9531 | -35.4648 | -34.1025
23.9592 | -35.4652 | -34.1088
23.9653 | -35.4657 | -34.115
23.9714 | -35.4661 | -34.1207
23.9775 | -35.4666 | -34.1264
23.9836 | -35.467 | -34.1324
23.9897 | -35.4674 | -34.1379
23.9958 | -35.4678 | -34.1435
24.002 | -35.4682 | -34.1496
24.0081 | -35.4686 | -34.1556
24.0142 | -35.469 | -34.1613
24.0203 | -35.4694 | -34.1672
24.0264 | -35.4699 | -34.173
24.0325 | -35.4703 | -34.1786
24.0386 | -35.4707 | -34.1842
24.0447 | -35.4711 | -34.1899
24.0508 | -35.4715 | -34.1953
24.0569 | -35.4719 | -34.2012
24.063 | -35.4724 | -34.2066
24.0691 | -35.4728 | -34.2118
24.0752 | -35.4732 | -34.2181
24.0813 | -35.4737 | -34.2241
24.0874 | -35.474 | -34.23
24.0935 | -35.4745 | -34.2357
24.0996 | -35.4748 | -34.2407
24.1057 | -35.4752 | -34.2461
24.1118 | -35.4757 | -34.2516
24.1179 | -35.4761 | -34.2574
24.124 | -35.4765 | -34.263
24.1301 | -35.477 | -34.2683
24.1362 | -35.4774 | -34.2739
24.1424 | -35.4778 | -34.2797
24.1485 | -35.4782 | -34.2856
24.1546 | -35.4785 | -34.2911
24.1607 | -35.4789 | -34.297
24.1668 | -35.4793 | -34.3026
24.1729 | -35.4797 | -34.3083
24.179 | -35.4801 | -34.3137
24.1851 | -35.4805 | -34.3196
24.1912 | -35.4809 | -34.3251
24.1973 | -35.4813 | -34.3308
24.2034 | -35.4817 | -34.3365
24.2095 | -35.4821 | -34.3426
24.2156 | -35.4825 | -34.3485
24.2217 | -35.483 | -34.3541
24.2278 | -35.4834 | -34.3599
24.2339 | -35.4838 | -34.3656
24.24 | -35.4841 | -34.371
24.2461 | -35.4845 | -34.3764
24.2522 | -35.4849 | -34.3823
24.2583 | -35.4854 | -34.3878
24.2644 | -35.4858 | -34.3935
24.2705 | -35.4862 | -34.3994
24.2766 | -35.4865 | -34.4051
24.2827 | -35.4869 | -34.4107
24.2889 | -35.4873 | -34.4164
24.295 | -35.4877 | -34.4219
24.3011 | -35.488 | -34.4277
24.3072 | -35.4884 | -34.4328
24.3133 | -35.4888 | -34.4386
24.3194 | -35.4892 | -34.4442
24.3255 | -35.4897 | -34.4497
24.3316 | -35.4901 | -34.4554
24.3377 | -35.4905 | -34.4609
24.3438 | -35.4909 | -34.4664
24.3499 | -35.4913 | -34.472
24.356 | -35.4916 | -34.4777
24.3621 | -35.4921 | -34.4835
24.3682 | -35.4925 | -34.4892
24.3743 | -35.4929 | -34.4946
24.3804 | -35.4933 | -34.4998
24.3865 | -35.4937 | -34.5057
24.3926 | -35.4941 | -34.5114
24.3987 | -35.4945 | -34.5172
24.4048 | -35.4949 | -34.5228
24.4109 | -35.4953 | -34.5278
24.417 | -35.4957 | -34.5337
24.4231 | -35.4961 | -34.5394
24.4293 | -35.4965 | -34.545
24.4354 | -35.4969 | -34.5504
24.4415 | -35.4973 | -34.5559
24.4476 | -35.4978 | -34.5617
24.4537 | -35.4982 | -34.5674
24.4598 | -35.4986 | -34.5731
24.4659 | -35.499 | -34.579
24.472 | -35.4995 | -34.5851
24.4781 | -35.4998 | -34.5907
24.4842 | -35.5002 | -34.597
24.4903 | -35.5006 | -34.6024
24.4964 | -35.5009 | -34.608
24.5025 | -35.5013 | -34.6137
24.5086 | -35.5017 | -34.6195
24.5147 | -35.5021 | -34.625
24.5208 | -35.5026 | -34.6306
24.5269 | -35.503 | -34.636
24.533 | -35.5034 | -34.6415
24.5391 | -35.5038 | -34.6472
24.5452 | -35.5042 | -34.653
24.5513 | -35.5046 | -34.6586
24.5574 | -35.505 | -34.6645
24.5635 | -35.5054 | -34.6704
24.5696 | -35.5058 | -34.6763
24.5758 | -35.5062 | -34.6818
24.5819 | -35.5066 | -34.6877
24.588 | -35.5069 | -34.693
24.5941 | -35.5074 | -34.6985
24.6002 | -35.5077 | -34.7041
24.6063 | -35.5081 | -34.7097
24.6124 | -35.5085 | -34.7154
24.6185 | -35.5089 | -34.721
24.6246 | -35.5092 | -34.7269
24.6307 | -35.5096 | -34.7325
24.6368 | -35.51 | -34.7384
24.6429 | -35.5104 | -34.744
24.649 | -35.5108 | -34.7495
24.6551 | -35.5112 | -34.7548
24.6612 | -35.5116 | -34.7603
24.6673 | -35.5121 | -34.7662
24.6734 | -35.5125 | -34.772
24.6795 | -35.5128 | -34.7779
24.6856 | -35.5132 | -34.7838
24.6917 | -35.5136 | -34.7896
24.6978 | -35.514 | -34.795
24.7039 | -35.5144 | -34.8006
24.71 | -35.5148 | -34.8063
24.7162 | -35.5153 | -34.8121
24.7223 | -35.5157 | -34.8178
24.7284 | -35.5161 | -34.8236
24.7345 | -35.5165 | -34.8295
24.7406 | -35.5169 | -34.8352
24.7467 | -35.5173 | -34.8411
24.7528 | -35.5177 | -34.8467
24.7589 | -35.5181 | -34.8522
24.765 | -35.5185 | -34.8579
24.7711 | -35.5189 | -34.8637
24.7772 | -35.5193 | -34.8693
24.7833 | -35.5197 | -34.875
24.7894 | -35.5201 | -34.8806
24.7955 | -35.5205 | -34.8862
24.8016 | -35.5208 | -34.8917
24.8077 | -35.5212 | -34.8975
24.8138 | -35.5216 | -34.903
24.8199 | -35.5221 | -34.9088
24.826 | -35.5224 | -34.9144
24.8321 | -35.5228 | -34.92
24.8382 | -35.5232 | -34.9257
24.8443 | -35.5236 | -34.9314
24.8504 | -35.524 | -34.9369
24.8565 | -35.5243 | -34.9424
24.8627 | -35.5247 | -34.9484
24.8688 | -35.5251 | -34.954
24.8749 | -35.5255 | -34.9593
24.881 | -35.5259 | -34.9648
24.8871 | -35.5262 | -34.9706
24.8932 | -35.5266 | -34.9763
24.8993 | -35.5271 | -34.9815
24.9054 | -35.5275 | -34.9872
24.9115 | -35.5279 | -34.9929
24.9176 | -35.5282 | -34.9982
24.9237 | -35.5286 | -35.004
24.9298 | -35.529 | -35.0093
24.9359 | -35.5293 | -35.0155
24.942 | -35.5297 | -35.0211
24.9481 | -35.5301 | -35.027
24.9542 | -35.5305 | -35.0324
24.9603 | -35.5309 | -35.0379
24.9664 | -35.5313 | -35.0433
24.9725 | -35.5317 | -35.049
24.9786 | -35.5321 | -35.0542
24.9847 | -35.5325 | -35.0597
24.9908 | -35.5329 | -35.0651
24.9969 | -35.5333 | -35.0708
25.0031 | -35.5337 | -35.0764
25.0092 | -35.5341 | -35.0818
25.0153 | -35.5344 | -35.0875
25.0214 | -35.5348 | -35.0927
25.0275 | -35.5351 | -35.0984
25.0336 | -35.5356 | -35.1039
25.0397 | -35.5359 | -35.1094
25.0458 | -35.5364 | -35.1146
25.0519 | -35.5368 | -35.1203
25.058 | -35.5371 | -35.1262
25.0641 | -35.5375 | -35.1317
25.0702 | -35.5379 | -35.1376
25.0763 | -35.5383 | -35.143
25.0824 | -35.5387 | -35.1483
25.0885 | -35.5392 | -35.154
25.0946 | -35.5395 | -35.1596
25.1007 | -35.5399 | -35.1653
25.1068 | -35.5403 | -35.1708
25.1129 | -35.5407 | -35.1764
25.119 | -35.541 | -35.1816
25.1251 | -35.5414 | -35.1872
25.1312 | -35.5418 | -35.1926
25.1373 | -35.5422 | -35.1982
25.1435 | -35.5426 | -35.2036
25.1496 | -35.543 | -35.2094
25.1557 | -35.5435 | -35.2151
25.1618 | -35.5439 | -35.2205
25.1679 | -35.5443 | -35.2264
25.174 | -35.5447 | -35.232
25.1801 | -35.5451 | -35.2377
25.1862 | -35.5454 | -35.2433
25.1923 | -35.5458 | -35.2489
25.1984 | -35.5462 | -35.2544
25.2045 | -35.5466 | -35.2599
25.2106 | -35.5469 | -35.2655
25.2167 | -35.5473 | -35.2709
25.2228 | -35.5477 | -35.2764
25.2289 | -35.548 | -35.2821
25.235 | -35.5485 | -35.2879
25.2411 | -35.5488 | -35.2933
25.2472 | -35.5492 | -35.2989
25.2533 | -35.5496 | -35.3043
25.2594 | -35.55 | -35.3097
25.2655 | -35.5503 | -35.3156
25.2716 | -35.5507 | -35.3209
25.2777 | -35.5511 | -35.3263
25.2838 | -35.5515 | -35.3319
25.29 | -35.5519 | -35.3376
25.2961 | -35.5523 | -35.3431
25.3022 | -35.5527 | -35.3483
25.3083 | -35.5531 | -35.3536
25.3144 | -35.5535 | -35.3593
25.3205 | -35.5538 | -35.365
25.3266 | -35.5542 | -35.3707
25.3327 | -35.5546 | -35.3762
25.3388 | -35.5549 | -35.3818
25.3449 | -35.5553 | -35.3872
25.351 | -35.5557 | -35.3926
25.3571 | -35.556 | -35.3983
25.3632 | -35.5563 | -35.4038
25.3693 | -35.5567 | -35.4094
25.3754 | -35.5572 | -35.4146
25.3815 | -35.5575 | -35.4201
25.3876 | -35.5579 | -35.4254
25.3937 | -35.5583 | -35.431
25.3998 | -35.5586 | -35.4365
25.4059 | -35.559 | -35.4424
25.412 | -35.5593 | -35.4478
25.4181 | -35.5597 | -35.4528
25.4242 | -35.5602 | -35.4585
25.4304 | -35.5605 | -35.4642
25.4365 | -35.5609 | -35.4699
25.4426 | -35.5612 | -35.4755
25.4487 | -35.5616 | -35.481
25.4548 | -35.562 | -35.4869
25.4609 | -35.5623 | -35.4927
25.467 | -35.5627 | -35.498
25.4731 | -35.5631 | -35.5035
25.4792 | -35.5634 | -35.5089
25.4853 | -35.5638 | -35.5146
25.4914 | -35.5642 | -35.5201
25.4975 | -35.5646 | -35.5255
25.5036 | -35.565 | -35.5308
25.5097 | -35.5653 | -35.5363
25.5158 | -35.5657 | -35.542
25.5219 | -35.5661 | -35.5477
25.528 | -35.5665 | -35.5534
25.5341 | -35.5669 | -35.5589
25.5402 | -35.5672 | -35.5645
25.5463 | -35.5676 | -35.5699
25.5524 | -35.568 | -35.5756
25.5585 | -35.5684 | -35.5815
25.5646 | -35.5688 | -35.5871
25.5707 | -35.5692 | -35.5926
25.5769 | -35.5696 | -35.5979
25.583 | -35.5699 | -35.6036
25.5891 | -35.5703 | -35.6092
25.5952 | -35.5707 | -35.615
25.6013 | -35.571 | -35.6204
25.6074 | -35.5714 | -35.6256
25.6135 | -35.5718 | -35.6308
25.6196 | -35.5722 | -35.6365
25.6257 | -35.5725 | -35.6421
25.6318 | -35.5729 | -35.6475
25.6379 | -35.5733 | -35.6531
25.644 | -35.5736 | -35.6582
25.6501 | -35.574 | -35.6637
25.6562 | -35.5744 | -35.6693
25.6623 | -35.5748 | -35.6747
25.6684 | -35.5752 | -35.6803
25.6745 | -35.5756 | -35.6856
25.6806 | -35.5759 | -35.6909
25.6867 | -35.5763 | -35.6961
25.6928 | -35.5767 | -35.7018
25.6989 | -35.577 | -35.7074
25.705 | -35.5774 | -35.7131
25.7111 | -35.5778 | -35.7187
25.7173 | -35.5782 | -35.7239
25.7234 | -35.5786 | -35.7295
25.7295 | -35.5789 | -35.7349
25.7356 | -35.5793 | -35.7405
25.7417 | -35.5796 | -35.7459
25.7478 | -35.5799 | -35.7512
25.7539 | -35.5803 | -35.7567
25.76 | -35.5807 | -35.7621
25.7661 | -35.5811 | -35.7677
25.7722 | -35.5814 | -35.7733
25.7783 | -35.5818 | -35.7786
25.7844 | -35.5821 | -35.7841
25.7905 | -35.5825 | -35.7896
25.7966 | -35.5829 | -35.7947
25.8027 | -35.5832 | -35.8001
25.8088 | -35.5836 | -35.8056
25.8149 | -35.584 | -35.8112
25.821 | -35.5842 | -35.8168
25.8271 | -35.5846 | -35.8225
25.8332 | -35.5851 | -35.828
25.8393 | -35.5854 | -35.8333
25.8454 | -35.5858 | -35.8389
25.8515 | -35.5862 | -35.8445
25.8576 | -35.5866 | -35.8496
25.8638 | -35.587 | -35.8552
25.8699 | -35.5873 | -35.8607
25.876 | -35.5877 | -35.8663
25.8821 | -35.5881 | -35.8717
25.8882 | -35.5885 | -35.8773
25.8943 | -35.5889 | -35.8829
25.9004 | -35.5892 | -35.8883
25.9065 | -35.5896 | -35.894
25.9126 | -35.59 | -35.8997
25.9187 | -35.5903 | -35.9052
25.9248 | -35.5907 | -35.9107
25.9309 | -35.5911 | -35.9165
25.937 | -35.5914 | -35.9221
25.9431 | -35.5918 | -35.9277
25.9492 | -35.5921 | -35.9331
25.9553 | -35.5925 | -35.9387
25.9614 | -35.5928 | -35.9443
25.9675 | -35.5932 | -35.9498
25.9736 | -35.5936 | -35.9555
25.9797 | -35.594 | -35.9614
25.9858 | -35.5943 | -35.9669
25.9919 | -35.5947 | -35.9727
25.998 | -35.595 | -35.9781
26.0042 | -35.5953 | -35.9835
26.0103 | -35.5956 | -35.989
26.0164 | -35.596 | -35.9944
26.0225 | -35.5964 | -36
26.0286 | -35.5968 | -36.0055
26.0347 | -35.5971 | -36.0108
26.0408 | -35.5975 | -36.0162
26.0469 | -35.5979 | -36.0217
26.053 | -35.5983 | -36.0272
26.0591 | -35.5986 | -36.0329
26.0652 | -35.599 | -36.038
26.0713 | -35.5994 | -36.0436
26.0774 | -35.5997 | -36.0489
26.0835 | -35.6001 | -36.0544
26.0896 | -35.6005 | -36.0598
26.0957 | -35.6009 | -36.0651
26.1018 | -35.6012 | -36.0701
26.1079 | -35.6016 | -36.0754
26.114 | -35.602 | -36.0806
26.1201 | -35.6023 | -36.0859
26.1262 | -35.6027 | -36.0913
26.1323 | -35.603 | -36.0965
26.1384 | -35.6034 | -36.1015
26.1445 | -35.6037 | -36.1072
26.1507 | -35.6041 | -36.1128
26.1568 | -35.6045 | -36.1183
26.1629 | -35.6049 | -36.1241
26.169 | -35.6052 | -36.1293
26.1751 | -35.6056 | -36.1347
26.1812 | -35.606 | -36.14
26.1873 | -35.6064 | -36.1453
26.1934 | -35.6068 | -36.1502
26.1995 | -35.6071 | -36.1558
26.2056 | -35.6074 | -36.1607
26.2117 | -35.6078 | -36.1664
26.2178 | -35.6082 | -36.1717
26.2239 | -35.6085 | -36.1769
26.23 | -35.6089 | -36.1824
26.2361 | -35.6092 | -36.1879
26.2422 | -35.6095 | -36.1931
26.2483 | -35.6099 | -36.1987
26.2544 | -35.6102 | -36.2044
26.2605 | -35.6106 | -36.2097
26.2666 | -35.611 | -36.2153
26.2727 | -35.6113 | -36.2208
26.2788 | -35.6117 | -36.2261
26.2849 | -35.612 | -36.2317
26.2911 | -35.6125 | -36.2369
26.2972 | -35.6128 | -36.2424
26.3033 | -35.6132 | -36.2477
26.3094 | -35.6135 | -36.2532
26.3155 | -35.6138 | -36.2579
26.3216 | -35.6142 | -36.2631
26.3277 | -35.6145 | -36.2681
26.3338 | -35.6149 | -36.2732
26.3399 | -35.6152 | -36.2784
26.346 | -35.6155 | -36.2837
26.3521 | -35.6159 | -36.2893
26.3582 | -35.6163 | -36.2946
26.3643 | -35.6167 | -36.3001
26.3704 | -35.6171 | -36.3056
26.3765 | -35.6175 | -36.3112
26.3826 | -35.6179 | -36.3161
26.3887 | -35.6182 | -36.3216
26.3948 | -35.6186 | -36.3269
26.4009 | -35.6189 | -36.3324
26.407 | -35.6192 | -36.3376
26.4131 | -35.6195 | -36.3435
26.4192 | -35.6198 | -36.3488
26.4253 | -35.6202 | -36.354
26.4314 | -35.6205 | -36.3595
26.4376 | -35.6209 | -36.3647
26.4437 | -35.6212 | -36.3701
26.4498 | -35.6215 | -36.3755
26.4559 | -35.6219 | -36.3806
26.462 | -35.6223 | -36.3855
26.4681 | -35.6226 | -36.3903
26.4742 | -35.623 | -36.3951
26.4803 | -35.6234 | -36.4004
26.4864 | -35.6237 | -36.4057
26.4925 | -35.624 | -36.4107
26.4986 | -35.6244 | -36.4162
26.5047 | -35.6248 | -36.4218
26.5108 | -35.6251 | -36.4268
26.5169 | -35.6255 | -36.4323
26.523 | -35.6259 | -36.4371
26.5291 | -35.6262 | -36.4424
26.5352 | -35.6265 | -36.4473
26.5413 | -35.6268 | -36.4525
26.5474 | -35.6272 | -36.4578
26.5535 | -35.6275 | -36.4635
26.5596 | -35.6278 | -36.4687
26.5657 | -35.6281 | -36.4738
26.5718 | -35.6285 | -36.4791
26.578 | -35.6289 | -36.4845
26.5841 | -35.6292 | -36.49
26.5902 | -35.6296 | -36.4956
26.5963 | -35.6299 | -36.5005
26.6024 | -35.6302 | -36.5056
26.6085 | -35.6305 | -36.5108
26.6146 | -35.6309 | -36.5162
26.6207 | -35.6311 | -36.5215
26.6268 | -35.6315 | -36.5269
26.6329 | -35.6317 | -36.5317
26.639 | -35.6321 | -36.537
26.6451 | -35.6325 | -36.5426
26.6512 | -35.6327 | -36.5478
26.6573 | -35.6331 | -36.553
26.6634 | -35.6334 | -36.5579
26.6695 | -35.6338 | -36.5629
26.6756 | -35.6342 | -36.5685
26.6817 | -35.6345 | -36.574
26.6878 | -35.6349 | -36.5795
26.6939 | -35.6353 | -36.5847
26.7 | -35.6356 | -36.5906
26.7061 | -35.6359 | -36.5956
26.7122 | -35.6363 | -36.6007
26.7183 | -35.6367 | -36.6059
26.7245 | -35.637 | -36.6113
26.7306 | -35.6373 | -36.6165
26.7367 | -35.6376 | -36.6215
26.7428 | -35.6379 | -36.6268
26.7489 | -35.6383 | -36.6319
26.755 | -35.6387 | -36.637
26.7611 | -35.6391 | -36.6423
26.7672 | -35.6395 | -36.648
26.7733 | -35.6398 | -36.6527
26.7794 | -35.6402 | -36.6581
26.7855 | -35.6405 | -36.6636
26.7916 | -35.6409 | -36.6689
26.7977 | -35.6412 | -36.6745
26.8038 | -35.6415 | -36.6799
26.8099 | -35.6418 | -36.6849
26.816 | -35.6422 | -36.6906
26.8221 | -35.6425 | -36.6955
26.8282 | -35.6429 | -36.7007
26.8343 | -35.6432 | -36.7064
26.8404 | -35.6436 | -36.7117
26.8465 | -35.6439 | -36.7175
26.8526 | -35.6443 | -36.7227
26.8587 | -35.6446 | -36.7281
26.8649 | -35.6449 | -36.7332
26.871 | -35.6453 | -36.7381
26.8771 | -35.6456 | -36.7429
26.8832 | -35.646 | -36.7483
26.8893 | -35.6464 | -36.7532
26.8954 | -35.6467 | -36.7587
26.9015 | -35.647 | -36.7643
26.9076 | -35.6473 | -36.7692
26.9137 | -35.6477 | -36.7745
26.9198 | -35.6479 | -36.7799
26.9259 | -35.6483 | -36.7851
26.932 | -35.6487 | -36.7904
26.9381 | -35.649 | -36.796
26.9442 | -35.6494 | -36.8005
26.9503 | -35.6498 | -36.8062
26.9564 | -35.6501 | -36.8112
26.9625 | -35.6504 | -36.8167
26.9686 | -35.6507 | -36.8221
26.9747 | -35.651 | -36.8277
26.9808 | -35.6514 | -36.833
26.9869 | -35.6518 | -36.838
26.993 | -35.6521 | -36.8432
26.9991 | -35.6524 | -36.8481
27.0052 | -35.6528 | -36.8535
27.0114 | -35.6531 | -36.8588
27.0175 | -35.6533 | -36.8642
27.0236 | -35.6537 | -36.8695
27.0297 | -35.6541 | -36.8748
27.0358 | -35.6544 | -36.8797
27.0419 | -35.6548 | -36.8848
27.048 | -35.6551 | -36.8903
27.0541 | -35.6554 | -36.8955
27.0602 | -35.6557 | -36.9009
27.0663 | -35.6561 | -36.9058
27.0724 | -35.6564 | -36.9116
27.0785 | -35.6568 | -36.9169
27.0846 | -35.6571 | -36.9217
27.0907 | -35.6575 | -36.9273
27.0968 | -35.6577 | -36.9325
27.1029 | -35.6581 | -36.9376
27.109 | -35.6584 | -36.9427
27.1151 | -35.6588 | -36.9485
27.1212 | -35.6591 | -36.9533
27.1273 | -35.6594 | -36.9587
27.1334 | -35.6598 | -36.9642
27.1395 | -35.6602 | -36.9693
27.1456 | -35.6605 | -36.9743
27.1518 | -35.6609 | -36.9797
27.1579 | -35.6612 | -36.9845
27.164 | -35.6615 | -36.9901
27.1701 | -35.6619 | -36.9956
27.1762 | -35.6622 | -37.0008
27.1823 | -35.6626 | -37.0058
27.1884 | -35.663 | -37.011
27.1945 | -35.6633 | -37.0158
27.2006 | -35.6637 | -37.0211
27.2067 | -35.6641 | -37.0264
27.2128 | -35.6645 | -37.0316
27.2189 | -35.6648 | -37.0372
27.225 | -35.6651 | -37.0422
27.2311 | -35.6654 | -37.0472
27.2372 | -35.6657 | -37.0526
27.2433 | -35.6662 | -37.0578
27.2494 | -35.6665 | -37.063
27.2555 | -35.6668 | -37.0689
27.2616 | -35.6671 | -37.0743
27.2677 | -35.6674 | -37.0796
27.2738 | -35.6678 | -37.0848
27.2799 | -35.6681 | -37.09
27.286 | -35.6685 | -37.0947
27.2921 | -35.6688 | -37.1004
27.2983 | -35.6691 | -37.1056
27.3044 | -35.6696 | -37.1105
27.3105 | -35.6699 | -37.116
27.3166 | -35.6703 | -37.121
27.3227 | -35.6706 | -37.1262
27.3288 | -35.671 | -37.1317
27.3349 | -35.6713 | -37.1375
27.341 | -35.6717 | -37.1421
27.3471 | -35.6721 | -37.1472
27.3532 | -35.6724 | -37.1517
27.3593 | -35.6728 | -37.1577
27.3654 | -35.6731 | -37.1631
27.3715 | -35.6735 | -37.1684
27.3776 | -35.6738 | -37.1738
27.3837 | -35.6741 | -37.1789
27.3898 | -35.6745 | -37.1839
27.3959 | -35.6748 | -37.1887
27.402 | -35.6752 | -37.1933
27.4081 | -35.6754 | -37.1985
27.4142 | -35.6757 | -37.2037
27.4203 | -35.6761 | -37.209
27.4264 | -35.6764 | -37.2141
27.4325 | -35.6767 | -37.2188
27.4387 | -35.6771 | -37.2242
27.4448 | -35.6774 | -37.2295
27.4509 | -35.6777 | -37.2351
27.457 | -35.678 | -37.2409
27.4631 | -35.6784 | -37.2465
27.4692 | -35.6787 | -37.2516
27.4753 | -35.6791 | -37.2564
27.4814 | -35.6794 | -37.2615
27.4875 | -35.6796 | -37.2663
27.4936 | -35.68 | -37.271
27.4997 | -35.6804 | -37.276
27.5058 | -35.6806 | -37.2817
27.5119 | -35.681 | -37.2865
27.518 | -35.6812 | -37.2918
27.5241 | -35.6814 | -37.2968
27.5302 | -35.6818 | -37.302
27.5363 | -35.6823 | -37.3071
27.5424 | -35.6826 | -37.3118
27.5485 | -35.6829 | -37.3173
27.5546 | -35.6832 | -37.3225
27.5607 | -35.6836 | -37.3274
27.5668 | -35.684 | -37.3321
27.5729 | -35.6843 | -37.3373
27.5791 | -35.6846 | -37.3418
27.5852 | -35.6848 | -37.3475
27.5913 | -35.6851 | -37.3528
27.5974 | -35.6853 | -37.3579
27.6035 | -35.6856 | -37.3631
27.6096 | -35.686 | -37.3676
27.6157 | -35.6863 | -37.3733
27.6218 | -35.6867 | -37.3784
27.6279 | -35.687 | -37.383
27.634 | -35.6875 | -37.3878
27.6401 | -35.6878 | -37.3932
27.6462 | -35.6881 | -37.3984
27.6523 | -35.6885 | -37.4037
27.6584 | -35.6888 | -37.4083
27.6645 | -35.6891 | -37.4134
27.6706 | -35.6894 | -37.4184
27.6767 | -35.6899 | -37.4241
27.6828 | -35.6901 | -37.4297
27.6889 | -35.6904 | -37.4354
27.695 | -35.6906 | -37.4401
27.7011 | -35.691 | -37.4448
27.7072 | -35.6913 | -37.4498
27.7133 | -35.6917 | -37.4552
27.7194 | -35.692 | -37.4604
27.7256 | -35.6922 | -37.4663
27.7317 | -35.6925 | -37.4715
27.7378 | -35.6928 | -37.4767
27.7439 | -35.6931 | -37.4818
27.75 | -35.6936 | -37.4873
27.7561 | -35.6938 | -37.4917
27.7622 | -35.6942 | -37.4964
27.7683 | -35.6946 | -37.5021
27.7744 | -35.6949 | -37.5071
27.7805 | -35.6951 | -37.5125
27.7866 | -35.6954 | -37.5173
27.7927 | -35.6957 | -37.5215
27.7988 | -35.696 | -37.5268
27.8049 | -35.6963 | -37.5314
27.811 | -35.6967 | -37.5362
27.8171 | -35.697 | -37.5416
27.8232 | -35.6973 | -37.5465
27.8293 | -35.6976 | -37.5516
27.8354 | -35.698 | -37.5573
27.8415 | -35.6983 | -37.5622
27.8476 | -35.6985 | -37.5669
27.8537 | -35.6988 | -37.5718
27.8598 | -35.6991 | -37.5774
27.866 | -35.6995 | -37.5825
27.8721 | -35.6999 | -37.588
27.8782 | -35.7002 | -37.5926
27.8843 | -35.7005 | -37.5982
27.8904 | -35.7008 | -37.6035
27.8965 | -35.7013 | -37.6077
27.9026 | -35.7016 | -37.6126
27.9087 | -35.702 | -37.618
27.9148 | -35.7023 | -37.6229
27.9209 | -35.7026 | -37.6277
27.927 | -35.7028 | -37.6329
27.9331 | -35.703 | -37.6382
27.9392 | -35.7034 | -37.6428
27.9453 | -35.7038 | -37.6485
27.9514 | -35.7041 | -37.6537
27.9575 | -35.7045 | -37.6589
27.9636 | -35.7047 | -37.6642
27.9697 | -35.7051 | -37.6692
27.9758 | -35.7053 | -37.6745
27.9819 | -35.7055 | -37.6801
27.988 | -35.7058 | -37.6853
27.9941 | -35.7061 | -37.6911
28.0002 | -35.7063 | -37.6955
28.0063 | -35.7067 | -37.7003
28.0125 | -35.707 | -37.7054
28.0186 | -35.7073 | -37.7109
28.0247 | -35.7076 | -37.7153
28.0308 | -35.708 | -37.7202
28.0369 | -35.7083 | -37.7261
28.043 | -35.7086 | -37.7306
28.0491 | -35.7089 | -37.7361
28.0552 | -35.7093 | -37.7412
28.0613 | -35.7095 | -37.7467
28.0674 | -35.7098 | -37.7519
28.0735 | -35.7099 | -37.7573

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.3313 | -41.9633
21.4908 | -35.5754 | -34.8478
16.4948 | -34.9328 | -27.5419
12.6602 | -34.5665 | -21.2849
9.7171 | -34.3432 | -16.4065
7.4582 | -34.1958 | -12.4403
5.7244 | -34.087 | -9.2187
4.3936 | -34.0217 | -6.6311
3.3722 | -34.0089 | -4.5856
2.5883 | -34.0084 | -3.037
1.9866 | -34.0008 | -1.7998
1.5248 | -34.0049 | -0.8076
1.1703 | -34.0271 | -0.16041
1 | -34.0174 | 0.15787
0.89822 | -34.0227 | 0.42658
0.68941 | -34.0108 | 1.0033
0.52919 | -34.0047 | 1.4514
0.40612 | -34.0042 | 2.0582
0.31171 | -34.0115 | 2.5001
0.23928 | -34.0069 | 3.223
0.18358 | -34.0079 | 3.642
0.14091 | -34.0383 | 4.5188
0.10819 | -34.0775 | 5.6729
0.083017 | -34.0827 | 7.1576
0.063705 | -34.1338 | 9.0239
0.048923 | -34.1923 | 11.5203
0.037479 | -34.303 | 14.6979
0.028801 | -34.4883 | 18.7166
0.022078 | -34.8048 | 24.0778
0.017023 | -35.2872 | 30.4712
0.013018 | -36.1006 | 38.1513
0.010014 | -37.2164 | 46.0522

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003727 | -7.206
0.007454 | -26.0769
0.011182 | -34.2172
0.014909 | -40.4213
0.018636 | -43.3922
0.022363 | -46.0816
0.02609 | -47.6281
0.029817 | -48.0618
0.033545 | -49.0777
0.037272 | -49.3926
0.040999 | -51.6532
0.044726 | -53.2997
0.048453 | -54.0497
0.05218 | -54.7387
0.055908 | -56.2093
0.059635 | -56.887
0.063362 | -57.5304
0.067089 | -57.3721
0.070816 | -57.5551
0.074543 | -58.7186
0.078271 | -59.9819
0.081998 | -61.4172
0.085725 | -60.3102
0.089452 | -60.3511
0.093179 | -61.8762
0.096906 | -61.0109
0.10063 | -61.4974
0.10436 | -63.3006
0.10809 | -64.1765
0.11181 | -64.4575
0.11554 | -64.4098
0.11927 | -63.966
0.123 | -64.4205
0.12672 | -64.9799
0.13045 | -65.8041
0.13418 | -66.2976
0.1379 | -65.8945
0.14163 | -66.3972
0.14536 | -67.3614
0.14909 | -67.6341
0.15281 | -69.1331
0.15654 | -67.7958
0.16027 | -67.9287
0.164 | -68.7343
0.16772 | -68.8093
0.17145 | -69.3182
0.17518 | -69.749
0.1789 | -70.1441
0.18263 | -70.4056
0.18636 | -69.8
0.19009 | -69.6904
0.19381 | -70.6395
0.19754 | -70.458
0.20127 | -71.2411
0.20499 | -71.1665
0.20872 | -71.0367
0.21245 | -71.7817
0.21618 | -71.6864
0.2199 | -71.3008
0.22363 | -71.8809
0.22736 | -72.9139
0.23109 | -73.0199
0.23481 | -72.9045
0.23854 | -72.8663
0.24227 | -73.8783
0.24599 | -73.8
0.24972 | -72.5237
0.25345 | -73.1635
0.25717 | -73.2112
0.2609 | -73.7644
0.26463 | -73.5891
0.26836 | -75.2452
0.27208 | -74.7847
0.27581 | -72.891
0.27954 | -72.2529
0.28326 | -72.596
0.28699 | -73.5213
0.29072 | -74.6895
0.29445 | -74.8205
0.29817 | -75.1032
0.3019 | -74.6453
0.30563 | -75.2277
0.30935 | -76.3646
0.31308 | -75.5908
0.31681 | -74.8553
0.32054 | -75.5136
0.32426 | -75.5449
0.32799 | -75.8955
0.33172 | -76.1965
0.33544 | -76.3947
0.33917 | -76.4143
0.3429 | -75.7318
0.34663 | -75.682
0.35035 | -76.6033
0.35408 | -77.5663
0.35781 | -77.2979
0.36154 | -76.4342
0.36526 | -77.4811
0.36899 | -77.3841
0.37272 | -76.8013
0.37644 | -76.9024
0.38017 | -77.7848
0.3839 | -77.7124
0.38763 | -77.4091
0.39135 | -76.2072
0.39508 | -77.5019
0.39881 | -78.2133
0.40253 | -78.0775
0.40626 | -77.8163
0.40999 | -78.2514
0.41372 | -77.4684
0.41744 | -77.8113
0.42117 | -77.4004
0.4249 | -77.683
0.42862 | -77.3682
0.43235 | -78.0184
0.43608 | -78.0298
0.43981 | -78.4444
0.44353 | -77.9594
0.44726 | -77.6494
0.45099 | -77.6921
0.45471 | -78.068
0.45844 | -79.3529
0.46217 | -79.9257
0.4659 | -79.9537
0.46962 | -79.2954
0.47335 | -78.7622
0.47708 | -79.162
0.4808 | -79.2304
0.48453 | -79.5058
0.48826 | -78.0297
0.49199 | -79.178
0.49571 | -80.9296
0.49944 | -80.2026
0.50317 | -79.5316
0.50689 | -79.6369
0.51062 | -80.0992
0.51435 | -80.834
0.51808 | -79.5564
0.5218 | -79.7399
0.52553 | -79.0923
0.52926 | -78.824
0.53299 | -80.2085
0.53671 | -80.8917
0.54044 | -80.9776
0.54417 | -81.1969
0.54789 | -81.2094
0.55162 | -80.9001
0.55535 | -80.9802
0.55908 | -81.4915
0.5628 | -81.9994
0.56653 | -81.689
0.57026 | -81.1645
0.57398 | -80.8906
0.57771 | -80.9314
0.58144 | -81.3858
0.58517 | -81.5333
0.58889 | -81.3788
0.59262 | -81.3231
0.59635 | -81.3351
0.60008 | -81.5233
0.6038 | -82.0261
0.60753 | -82.2243
0.61126 | -82.0719
0.61498 | -81.9316
0.61871 | -82.0991
0.62244 | -82.1486
0.62616 | -82.2978
0.62989 | -82.3371
0.63362 | -81.6684
0.63735 | -81.1501
0.64107 | -81.0741
0.6448 | -81.1615
0.64853 | -81.6264
0.65226 | -82.2383
0.65598 | -82.8915
0.65971 | -83.2663
0.66344 | -82.8345
0.66716 | -82.3053
0.67089 | -82.1233
0.67462 | -82.2492
0.67834 | -82.8462
0.68207 | -83.1349
0.6858 | -82.9859
0.68953 | -83.0264
0.69325 | -83.3359
0.69698 | -83.7177
0.70071 | -83.8725
0.70444 | -83.8827
0.70816 | -83.6725
0.71189 | -83.4722
0.71562 | -83.3502
0.71934 | -83.1395
0.72307 | -82.9502
0.7268 | -82.9756
0.73053 | -83.2726
0.73425 | -83.7686
0.73798 | -84.4093
0.74171 | -84.3667
0.74543 | -83.8327
0.74916 | -83.6415
0.75289 | -83.3972
0.75662 | -83.0366
0.76034 | -82.9068
0.76407 | -83.4934
0.7678 | -84.2574
0.77152 | -84.2974
0.77525 | -83.8675
0.77898 | -83.6807
0.78271 | -83.9455
0.78643 | -84.4314
0.79016 | -84.7151
0.79389 | -84.5616
0.79761 | -84.6899
0.80134 | -84.9638
0.80507 | -84.6423
0.8088 | -84.2624
0.81252 | -84.3894
0.81625 | -84.79
0.81998 | -85.0127
0.82371 | -85.0552
0.82743 | -85.2272
0.83116 | -85.2943
0.83489 | -85.3031
0.83861 | -85.225
0.84234 | -85.0685
0.84607 | -84.7627
0.84979 | -84.5348
0.85352 | -84.5878
0.85725 | -84.7144
0.86098 | -84.4808
0.8647 | -84.18
0.86843 | -84.2731
0.87216 | -84.3936
0.87589 | -84.17
0.87961 | -84.0075
0.88334 | -84.3261
0.88707 | -84.6243
0.89079 | -84.5483
0.89452 | -84.4918
0.89825 | -84.9996
0.90197 | -85.6755
0.9057 | -85.8713
0.90943 | -85.7523
0.91316 | -85.5391
0.91688 | -85.5149
0.92061 | -85.6616
0.92434 | -86.3004
0.92807 | -86.7568
0.93179 | -86.3863
0.93552 | -86.038
0.93925 | -86.1596
0.94297 | -86.5023
0.9467 | -86.5158
0.95043 | -86.4771
0.95416 | -86.4583
0.95788 | -86.2319
0.96161 | -86.087
0.96534 | -86.3707
0.96906 | -86.4196
0.97279 | -85.9946
0.97652 | -85.6244
0.98025 | -85.4977
0.98397 | -85.5487
0.9877 | -85.8026
0.99143 | -85.8265
0.99516 | -85.1634
0.99888 | -84.7914
1.0026 | -85.5545
1.0063 | -86.8301
1.0101 | -87.1661
1.0138 | -86.891
1.0175 | -86.8151
1.0212 | -87.0971
1.025 | -87.4947
1.0287 | -87.1589
1.0324 | -86.2758
1.0362 | -85.8176
1.0399 | -85.989
1.0436 | -86.4011
1.0473 | -86.9544
1.0511 | -87.3567
1.0548 | -87.4038
1.0585 | -87.3593
1.0622 | -87.2429
1.066 | -86.9452
1.0697 | -86.1372
1.0734 | -85.9664
1.0772 | -86.1181
1.0809 | -86.5243
1.0846 | -86.9633
1.0883 | -87.4319
1.0921 | -87.8336
1.0958 | -87.9441
1.0995 | -87.8782
1.1032 | -87.7492
1.107 | -87.5091
1.1107 | -87.2873
1.1144 | -87.2747
1.1182 | -87.4611
1.1219 | -87.6477
1.1256 | -87.7572
1.1293 | -87.894
1.1331 | -87.9631
1.1368 | -87.744
1.1405 | -87.3525
1.1442 | -87.0734
1.148 | -87.0369
1.1517 | -87.2549
1.1554 | -87.5772
1.1591 | -87.7555
1.1629 | -87.6962
1.1666 | -87.4353
1.1703 | -87.2153
1.1741 | -87.2792
1.1778 | -87.5239
1.1815 | -87.7105
1.1852 | -87.8234
1.189 | -88.0286
1.1927 | -88.2513
1.1964 | -88.1399
1.2001 | -87.8004
1.2039 | -87.6595
1.2076 | -87.8004
1.2113 | -87.9762
1.2151 | -87.9735
1.2188 | -87.9056
1.2225 | -87.9875
1.2262 | -88.264
1.23 | -88.5563
1.2337 | -88.6238
1.2374 | -88.4896
1.2411 | -88.3852
1.2449 | -88.4504
1.2486 | -88.5947
1.2523 | -88.6434
1.2561 | -88.578
1.2598 | -88.396
1.2635 | -88.2633
1.2672 | -88.3846
1.271 | -88.5492
1.2747 | -88.4353
1.2784 | -88.0347
1.2821 | -87.5822
1.2859 | -87.4629
1.2896 | -87.8214
1.2933 | -88.3455
1.2971 | -88.5949
1.3008 | -88.5361
1.3045 | -88.4126
1.3082 | -88.4194
1.312 | -88.656
1.3157 | -88.9998
1.3194 | -89.2279
1.3231 | -89.3574
1.3269 | -89.4968
1.3306 | -89.6658
1.3343 | -89.8138
1.3381 | -89.7908
1.3418 | -89.447
1.3455 | -88.937
1.3492 | -88.5509
1.353 | -88.4253
1.3567 | -88.5928
1.3604 | -88.9132
1.3641 | -89.0651
1.3679 | -89.0201
1.3716 | -88.9829
1.3753 | -89.0263
1.3791 | -89.1486
1.3828 | -89.2944
1.3865 | -89.3786
1.3902 | -89.3869
1.394 | -89.3379
1.3977 | -89.1326
1.4014 | -88.8653
1.4051 | -88.9212
1.4089 | -89.3331
1.4126 | -89.6698
1.4163 | -89.6288
1.4201 | -89.3564
1.4238 | -89.234
1.4275 | -89.4178
1.4312 | -89.6935
1.435 | -89.8451
1.4387 | -89.8551
1.4424 | -89.7508
1.4461 | -89.6345
1.4499 | -89.6784
1.4536 | -89.9946
1.4573 | -90.3998
1.4611 | -90.5039
1.4648 | -90.3128
1.4685 | -90.1638
1.4722 | -90.2199
1.476 | -90.3516
1.4797 | -90.2575
1.4834 | -89.9461
1.4871 | -89.7469
1.4909 | -89.8548
1.4946 | -90.1419
1.4983 | -90.3301
1.5021 | -90.376
1.5058 | -90.3795
1.5095 | -90.4123
1.5132 | -90.5217
1.517 | -90.5051
1.5207 | -90.1726
1.5244 | -89.79
1.5281 | -89.5496
1.5319 | -89.4895
1.5356 | -89.7201
1.5393 | -90.205
1.543 | -90.6557
1.5468 | -90.8946
1.5505 | -91.0262
1.5542 | -91.1578
1.558 | -91.2044
1.5617 | -91.058
1.5654 | -90.7549
1.5691 | -90.4937
1.5729 | -90.391
1.5766 | -90.2979
1.5803 | -90.1345
1.584 | -89.9875
1.5878 | -89.9292
1.5915 | -90.0576
1.5952 | -90.3329
1.599 | -90.4555
1.6027 | -90.3927
1.6064 | -90.4863
1.6101 | -90.8885
1.6139 | -91.161
1.6176 | -91.2692
1.6213 | -91.1685
1.625 | -90.9695
1.6288 | -90.824
1.6325 | -90.8024
1.6362 | -90.9003
1.64 | -91.07
1.6437 | -91.2359
1.6474 | -91.3521
1.6511 | -91.4199
1.6549 | -91.4391
1.6586 | -91.3756
1.6623 | -91.2425
1.666 | -91.1091
1.6698 | -91.0275
1.6735 | -90.9741
1.6772 | -90.9396
1.681 | -90.9583
1.6847 | -91.0562
1.6884 | -91.2121
1.6921 | -91.3779
1.6959 | -91.4821
1.6996 | -91.4726
1.7033 | -91.3395
1.707 | -91.1648
1.7108 | -91.0653
1.7145 | -91.131
1.7182 | -91.3736
1.722 | -91.7132
1.7257 | -91.9951
1.7294 | -92.0975
1.7331 | -92.0191
1.7369 | -91.8466
1.7406 | -91.6905
1.7443 | -91.5713
1.748 | -91.4646
1.7518 | -91.3542
1.7555 | -91.2671
1.7592 | -91.2302
1.763 | -91.2443
1.7667 | -91.3313
1.7704 | -91.5037
1.7741 | -91.6681
1.7779 | -91.7444
1.7816 | -91.7495
1.7853 | -91.7409
1.789 | -91.7155
1.7928 | -91.6757
1.7965 | -91.6764
1.8002 | -91.746
1.804 | -91.8252
1.8077 | -91.8569
1.8114 | -91.8404
1.8151 | -91.8282
1.8189 | -91.8708
1.8226 | -91.9644
1.8263 | -92.054
1.83 | -92.0872
1.8338 | -92.0755
1.8375 | -92.0822
1.8412 | -92.1216
1.845 | -92.1614
1.8487 | -92.1712
1.8524 | -92.1385
1.8561 | -92.0414
1.8599 | -91.8556
1.8636 | -91.6005
1.8673 | -91.3634
1.871 | -91.2434
1.8748 | -91.2663
1.8785 | -91.3822
1.8822 | -91.5099
1.8859 | -91.6191
1.8897 | -91.6836
1.8934 | -91.6818
1.8971 | -91.6631
1.9009 | -91.7294
1.9046 | -91.9204
1.9083 | -92.1503
1.912 | -92.2988
1.9158 | -92.3964
1.9195 | -92.5148
1.9232 | -92.6214
1.9269 | -92.5865
1.9307 | -92.3938
1.9344 | -92.1599
1.9381 | -91.9855
1.9419 | -91.9303
1.9456 | -92.0089
1.9493 | -92.1493
1.953 | -92.262
1.9568 | -92.268
1.9605 | -92.1668
1.9642 | -91.9776
1.9679 | -91.7476
1.9717 | -91.5285
1.9754 | -91.3793
1.9791 | -91.3704
1.9829 | -91.5132
1.9866 | -91.7326
1.9903 | -91.9257
1.994 | -91.9968
1.9978 | -91.9679
2.0015 | -91.9645
2.0052 | -92.0827
2.0089 | -92.2921
2.0127 | -92.4932
2.0164 | -92.6061
2.0201 | -92.6021
2.0239 | -92.4942
2.0276 | -92.3651
2.0313 | -92.32
2.035 | -92.4147
2.0388 | -92.6079
2.0425 | -92.7868
2.0462 | -92.898
2.0499 | -92.9798
2.0537 | -92.973
2.0574 | -92.8202
2.0611 | -92.6408
2.0649 | -92.6147
2.0686 | -92.7839
2.0723 | -93.0652
2.076 | -93.3128
2.0798 | -93.3339
2.0835 | -93.0686
2.0872 | -92.6688
2.0909 | -92.3225
2.0947 | -92.1084
2.0984 | -92.056
2.1021 | -92.1155
2.1059 | -92.2126
2.1096 | -92.3335
2.1133 | -92.4811
2.117 | -92.6192
2.1208 | -92.7446
2.1245 | -92.8808
2.1282 | -93.0397
2.1319 | -93.1992
2.1357 | -93.37
2.1394 | -93.5097
2.1431 | -93.5248
2.1469 | -93.3731
2.1506 | -93.0599
2.1543 | -92.8173
2.158 | -92.612
2.1618 | -92.4547
2.1655 | -92.3396
2.1692 | -92.251
2.1729 | -92.1875
2.1767 | -92.1604
2.1804 | -92.2004
2.1841 | -92.3339
2.1878 | -92.5615
2.1916 | -92.8464
2.1953 | -93.1126
2.199 | -93.2939
2.2028 | -93.3825
2.2065 | -93.3917
2.2102 | -93.3527
2.2139 | -93.287
2.2177 | -93.2251
2.2214 | -93.1867
2.2251 | -93.2007
2.2288 | -93.2737
2.2326 | -93.3761
2.2363 | -93.4607
2.24 | -93.489
2.2438 | -93.4218
2.2475 | -93.2712
2.2512 | -93.0854
2.2549 | -92.9447
2.2587 | -92.9096
2.2624 | -92.9966
2.2661 | -93.1916
2.2698 | -93.4579
2.2736 | -93.7592
2.2773 | -93.9987
2.281 | -94.1109
2.2848 | -94.0994
2.2885 | -94.0059
2.2922 | -93.8861
2.2959 | -93.7801
2.2997 | -93.7154
2.3034 | -93.7045
2.3071 | -93.7184
2.3108 | -93.6931
2.3146 | -93.5934
2.3183 | -93.4845
2.322 | -93.4253
2.3258 | -93.4689
2.3295 | -93.6221
2.3332 | -93.7967
2.3369 | -93.8988
2.3407 | -93.8819
2.3444 | -93.7598
2.3481 | -93.5851
2.3518 | -93.4055
2.3556 | -93.2388
2.3593 | -93.0912
2.363 | -92.9765
2.3668 | -92.9022
2.3705 | -92.8775
2.3742 | -92.9044
2.3779 | -92.9576
2.3817 | -93.0154
2.3854 | -93.0635
2.3891 | -93.1184
2.3928 | -93.2382
2.3966 | -93.4673
2.4003 | -93.7964
2.404 | -94.1531
2.4078 | -94.4639
2.4115 | -94.689
2.4152 | -94.845
2.4189 | -94.929
2.4227 | -94.957
2.4264 | -94.9582
2.4301 | -94.9481
2.4338 | -94.8911
2.4376 | -94.7425
2.4413 | -94.4993
2.445 | -94.2094
2.4488 | -93.9358
2.4525 | -93.758
2.4562 | -93.7205
2.4599 | -93.8223
2.4637 | -94.0143
2.4674 | -94.2161
2.4711 | -94.3476
2.4748 | -94.3742
2.4786 | -94.2893
2.4823 | -94.109
2.486 | -93.924
2.4897 | -93.8139
2.4935 | -93.7888
2.4972 | -93.8289
2.5009 | -93.9066
2.5047 | -94.0033
2.5084 | -94.1155
2.5121 | -94.2332
2.5158 | -94.3411
2.5196 | -94.4261
2.5233 | -94.4795
2.527 | -94.4674
2.5307 | -94.4064
2.5345 | -94.3296
2.5382 | -94.2933
2.5419 | -94.3552
2.5457 | -94.551
2.5494 | -94.8589
2.5531 | -95.1707
2.5568 | -95.3631
2.5606 | -95.384
2.5643 | -95.2528
2.568 | -95.0664
2.5717 | -94.9186
2.5755 | -94.8722
2.5792 | -94.9238
2.5829 | -95.0483
2.5867 | -95.1995
2.5904 | -95.3174
2.5941 | -95.3426
2.5978 | -95.2238
2.6016 | -94.9774
2.6053 | -94.6922
2.609 | -94.4518
2.6127 | -94.3258
2.6165 | -94.3397
2.6202 | -94.4818
2.6239 | -94.6893
2.6277 | -94.8715
2.6314 | -94.9439
2.6351 | -94.8676
2.6388 | -94.68
2.6426 | -94.474
2.6463 | -94.3331
2.65 | -94.2861
2.6537 | -94.342
2.6575 | -94.4754
2.6612 | -94.6298
2.6649 | -94.7518
2.6687 | -94.7817
2.6724 | -94.7098
2.6761 | -94.6013
2.6798 | -94.5247
2.6836 | -94.5231
2.6873 | -94.6544
2.691 | -94.7668
2.6947 | -94.9181
2.6985 | -95.0791
2.7022 | -95.2339
2.7059 | -95.3498
2.7097 | -95.4054
2.7134 | -95.4101
2.7171 | -95.4113
2.7208 | -95.4253
2.7246 | -95.4548
2.7283 | -95.4827
2.732 | -95.4848
2.7357 | -95.4425
2.7395 | -95.3651
2.7432 | -95.2776
2.7469 | -95.214
2.7507 | -95.2094
2.7544 | -95.2807
2.7581 | -95.3987
2.7618 | -95.5224
2.7656 | -95.5793
2.7693 | -95.5514
2.773 | -95.4535
2.7767 | -95.3183
2.7805 | -95.1954
2.7842 | -95.1167
2.7879 | -95.0997
2.7917 | -95.1369
2.7954 | -95.1991
2.7991 | -95.2676
2.8028 | -95.3303
2.8066 | -95.3866
2.8103 | -95.4453
2.814 | -95.529
2.8177 | -95.6314
2.8215 | -95.7448
2.8252 | -95.8534
2.8289 | -95.9383
2.8327 | -95.9751
2.8364 | -95.9473
2.8401 | -95.8645
2.8438 | -95.7524
2.8476 | -95.6376
2.8513 | -95.5567
2.855 | -95.506
2.8587 | -95.4622
2.8625 | -95.4046
2.8662 | -95.3483
2.8699 | -95.3245
2.8736 | -95.3629
2.8774 | -95.4624
2.8811 | -95.5908
2.8848 | -95.6919
2.8886 | -95.7224
2.8923 | -95.67
2.896 | -95.549
2.8997 | -95.4024
2.9035 | -95.2809
2.9072 | -95.2015
2.9109 | -95.1769
2.9146 | -95.2175
2.9184 | -95.3227
2.9221 | -95.4761
2.9258 | -95.6519
2.9296 | -95.7898
2.9333 | -95.8238
2.937 | -95.7319
2.9407 | -95.5606
2.9445 | -95.4025
2.9482 | -95.3316
2.9519 | -95.3879
2.9556 | -95.5642
2.9594 | -95.7945
2.9631 | -95.9915
2.9668 | -96.0851
2.9706 | -96.063
2.9743 | -95.9584
2.978 | -95.8296
2.9817 | -95.7248
2.9855 | -95.6616
2.9892 | -95.6167
2.9929 | -95.5655
2.9966 | -95.5084
3.0004 | -95.4399
3.0041 | -95.3668
3.0078 | -95.3068
3.0116 | -95.2607
3.0153 | -95.2495
3.019 | -95.2979
3.0227 | -95.3981
3.0265 | -95.5564
3.0302 | -95.7423
3.0339 | -95.9244
3.0376 | -96.0874
3.0414 | -96.1963
3.0451 | -96.2261
3.0488 | -96.1954
3.0526 | -96.1403
3.0563 | -96.0873
3.06 | -96.07
3.0637 | -96.0872
3.0675 | -96.1075
3.0712 | -96.1043
3.0749 | -96.0516
3.0786 | -95.9474
3.0824 | -95.8116
3.0861 | -95.6563
3.0898 | -95.5193
3.0936 | -95.4221
3.0973 | -95.3689
3.101 | -95.3578
3.1047 | -95.4041
3.1085 | -95.5141
3.1122 | -95.6812
3.1159 | -95.8781
3.1196 | -96.0444
3.1234 | -96.1175
3.1271 | -96.0869
3.1308 | -95.9667
3.1346 | -95.8183
3.1383 | -95.7061
3.142 | -95.6681
3.1457 | -95.7092
3.1495 | -95.8148
3.1532 | -95.9618
3.1569 | -96.1117
3.1606 | -96.2441
3.1644 | -96.3543
3.1681 | -96.4344
3.1718 | -96.47
3.1755 | -96.4659
3.1793 | -96.4477
3.183 | -96.4317
3.1867 | -96.4195
3.1905 | -96.4181
3.1942 | -96.4387
3.1979 | -96.4713
3.2016 | -96.5182
3.2054 | -96.5915
3.2091 | -96.6703
3.2128 | -96.7342
3.2165 | -96.7674
3.2203 | -96.7776
3.224 | -96.7822
3.2277 | -96.7952
3.2315 | -96.8035
3.2352 | -96.7927
3.2389 | -96.753
3.2426 | -96.681
3.2464 | -96.5686
3.2501 | -96.4398
3.2538 | -96.3114
3.2575 | -96.2101
3.2613 | -96.1539
3.265 | -96.1503
3.2687 | -96.2036
3.2725 | -96.3
3.2762 | -96.4327
3.2799 | -96.5771
3.2836 | -96.7069
3.2874 | -96.8034
3.2911 | -96.8773
3.2948 | -96.9066
3.2985 | -96.887
3.3023 | -96.8479
3.306 | -96.7767
3.3097 | -96.6818
3.3135 | -96.568
3.3172 | -96.4637
3.3209 | -96.3936
3.3246 | -96.3671
3.3284 | -96.3861
3.3321 | -96.4322
3.3358 | -96.4877
3.3395 | -96.5306
3.3433 | -96.5438
3.347 | -96.5173
3.3507 | -96.4777
3.3545 | -96.4293
3.3582 | -96.3585
3.3619 | -96.2861
3.3656 | -96.2172
3.3694 | -96.1707
3.3731 | -96.1359
3.3768 | -96.1284
3.3805 | -96.1645
3.3843 | -96.246
3.388 | -96.377
3.3917 | -96.5474
3.3955 | -96.7378
3.3992 | -96.9297
3.4029 | -97.092
3.4066 | -97.1912
3.4104 | -97.2247
3.4141 | -97.1915
3.4178 | -97.1037
3.4215 | -96.9826
3.4253 | -96.8532
3.429 | -96.755
3.4327 | -96.6881
3.4365 | -96.6692
3.4402 | -96.7009
3.4439 | -96.7773
3.4476 | -96.8928
3.4514 | -97.0355
3.4551 | -97.1856
3.4588 | -97.3226
3.4625 | -97.44
3.4663 | -97.5244
3.47 | -97.5738
3.4737 | -97.5918
3.4775 | -97.5808
3.4812 | -97.5439
3.4849 | -97.4846
3.4886 | -97.4106
3.4924 | -97.3254
3.4961 | -97.2626
3.4998 | -97.2304
3.5035 | -97.2322
3.5073 | -97.2627
3.511 | -97.3099
3.5147 | -97.3609
3.5184 | -97.3881
3.5222 | -97.3897
3.5259 | -97.3668
3.5296 | -97.3154
3.5334 | -97.2508
3.5371 | -97.1861
3.5408 | -97.1485
3.5445 | -97.1413
3.5483 | -97.175
3.552 | -97.24
3.5557 | -97.3311
3.5594 | -97.4229
3.5632 | -97.4912
3.5669 | -97.5347
3.5706 | -97.5642
3.5744 | -97.595
3.5781 | -97.6056
3.5818 | -97.6081
3.5855 | -97.5937
3.5893 | -97.5643
3.593 | -97.5315
3.5967 | -97.4968
3.6004 | -97.4812
3.6042 | -97.4693
3.6079 | -97.4504
3.6116 | -97.4264
3.6154 | -97.3957
3.6191 | -97.36
3.6228 | -97.331
3.6265 | -97.3253
3.6303 | -97.3519
3.634 | -97.4213
3.6377 | -97.4908
3.6414 | -97.5379
3.6452 | -97.5536
3.6489 | -97.5346
3.6526 | -97.5021
3.6564 | -97.4623
3.6601 | -97.4333
3.6638 | -97.4293
3.6675 | -97.4448
3.6713 | -97.4857
3.675 | -97.5449
3.6787 | -97.601
3.6824 | -97.6406
3.6862 | -97.6644
3.6899 | -97.6795
3.6936 | -97.7043
3.6974 | -97.7492
3.7011 | -97.7915
3.7048 | -97.8193
3.7085 | -97.8108
3.7123 | -97.7803
3.716 | -97.7539
3.7197 | -97.7388
3.7234 | -97.7638
3.7272 | -97.8222
3.7309 | -97.8907
3.7346 | -97.941
3.7384 | -97.946
3.7421 | -97.9
3.7458 | -97.8148
3.7495 | -97.6989
3.7533 | -97.5727
3.757 | -97.4738
3.7607 | -97.4421
3.7644 | -97.3964
3.7682 | -97.3719
3.7719 | -97.363
3.7756 | -97.3613
3.7794 | -97.3825
3.7831 | -97.4181
3.7868 | -97.4735
3.7905 | -97.5672
3.7943 | -97.6908
3.798 | -97.8352
3.8017 | -97.9911
3.8054 | -98.1343
3.8092 | -98.2588
3.8129 | -98.3487
3.8166 | -98.39
3.8203 | -98.3958
3.8241 | -98.3723
3.8278 | -98.3158
3.8315 | -98.243
3.8353 | -98.1631
3.839 | -98.0878
3.8427 | -98.0072
3.8464 | -97.9349
3.8502 | -97.8751
3.8539 | -97.8229
3.8576 | -97.7919
3.8613 | -97.7888
3.8651 | -97.8038
3.8688 | -97.8434
3.8725 | -97.8999
3.8763 | -97.9603
3.88 | -98.0238
3.8837 | -98.0893
3.8874 | -98.1374
3.8912 | -98.1563
3.8949 | -98.145
3.8986 | -98.1021
3.9023 | -98.0425
3.9061 | -97.9816
3.9098 | -97.9151
3.9135 | -97.8634
3.9173 | -97.8295
3.921 | -97.8144
3.9247 | -97.8064
3.9284 | -97.8162
3.9322 | -97.8434
3.9359 | -97.8805
3.9396 | -97.9352
3.9433 | -97.9992
3.9471 | -98.0664
3.9508 | -98.1404
3.9545 | -98.2062
3.9583 | -98.2424
3.962 | -98.2609
3.9657 | -98.2558
3.9694 | -98.2398
3.9732 | -98.2336
3.9769 | -98.2278
3.9806 | -98.2292
3.9843 | -98.2528
3.9881 | -98.3059
3.9918 | -98.3802
3.9955 | -98.4707
3.9993 | -98.5651
4.003 | -98.6435
4.0067 | -98.6874
4.0104 | -98.6858
4.0142 | -98.6225
4.0179 | -98.5148
4.0216 | -98.3654
4.0253 | -98.1954
4.0291 | -98.0452
4.0328 | -97.9238
4.0365 | -97.8339
4.0403 | -97.79
4.044 | -97.7941
4.0477 | -97.8227
4.0514 | -97.8681
4.0552 | -97.9176
4.0589 | -97.9622
4.0626 | -98.0084
4.0663 | -98.0408
4.0701 | -98.0628
4.0738 | -98.0842
4.0775 | -98.1227
4.0813 | -98.1726
4.085 | -98.2492
4.0887 | -98.3426
4.0924 | -98.4413
4.0962 | -98.5392
4.0999 | -98.6205
4.1036 | -98.6801
4.1073 | -98.7156
4.1111 | -98.7398
4.1148 | -98.7521
4.1185 | -98.7586
4.1223 | -98.7565
4.126 | -98.7321
4.1297 | -98.6961
4.1334 | -98.6438
4.1372 | -98.5683
4.1409 | -98.4743
4.1446 | -98.3744
4.1483 | -98.277
4.1521 | -98.2002
4.1558 | -98.1525
4.1595 | -98.1289
4.1632 | -98.1308
4.167 | -98.1323
4.1707 | -98.1277
4.1744 | -98.1211
4.1782 | -98.1107
4.1819 | -98.1028
4.1856 | -98.1111
4.1893 | -98.144
4.1931 | -98.2014
4.1968 | -98.2767
4.2005 | -98.3697
4.2042 | -98.475
4.208 | -98.5791
4.2117 | -98.6665
4.2154 | -98.7338
4.2192 | -98.7715
4.2229 | -98.7894
4.2266 | -98.7774
4.2303 | -98.7507
4.2341 | -98.7148
4.2378 | -98.6751
4.2415 | -98.6483
4.2452 | -98.6326
4.249 | -98.6365
4.2527 | -98.66
4.2564 | -98.6981
4.2602 | -98.7478
4.2639 | -98.7885
4.2676 | -98.8244
4.2713 | -98.8448
4.2751 | -98.8405
4.2788 | -98.8026
4.2825 | -98.7437
4.2862 | -98.6728
4.29 | -98.6021
4.2937 | -98.5404
4.2974 | -98.5081
4.3012 | -98.5118
4.3049 | -98.4918
4.3086 | -98.477
4.3123 | -98.4659
4.3161 | -98.4475
4.3198 | -98.4218
4.3235 | -98.3955
4.3272 | -98.3776
4.331 | -98.3733
4.3347 | -98.3739
4.3384 | -98.3849
4.3422 | -98.4063
4.3459 | -98.4358
4.3496 | -98.4598
4.3533 | -98.4722
4.3571 | -98.4839
4.3608 | -98.4929
4.3645 | -98.4985
4.3682 | -98.5144
4.372 | -98.5515
4.3757 | -98.6009
4.3794 | -98.6562
4.3832 | -98.7305
4.3869 | -98.8138
4.3906 | -98.895
4.3943 | -98.9709
4.3981 | -99.037
4.4018 | -99.1015
4.4055 | -99.1585
4.4092 | -99.212
4.413 | -99.2541
4.4167 | -99.2874
4.4204 | -99.3134
4.4242 | -99.3207
4.4279 | -99.3114
4.4316 | -99.2934
4.4353 | -99.2607
4.4391 | -99.2337
4.4428 | -99.2169
4.4465 | -99.2112
4.4502 | -99.2219
4.454 | -99.2376
4.4577 | -99.2732
4.4614 | -99.3209
4.4652 | -99.3593
4.4689 | -99.388
4.4726 | -99.4006
4.4763 | -99.4031
4.4801 | -99.3923
4.4838 | -99.3773
4.4875 | -99.3625
4.4912 | -99.3575
4.495 | -99.3651
4.4987 | -99.3741
4.5024 | -99.3878
4.5061 | -99.4039
4.5099 | -99.4094
4.5136 | -99.4072
4.5173 | -99.4069
4.5211 | -99.4117
4.5248 | -99.4353
4.5285 | -99.4829
4.5322 | -99.5528
4.536 | -99.6365
4.5397 | -99.7259
4.5434 | -99.8134
4.5471 | -99.8905
4.5509 | -99.9535
4.5546 | -99.9926
4.5583 | -100.0022
4.5621 | -99.999
4.5658 | -99.9737
4.5695 | -99.9238
4.5732 | -99.8478
4.577 | -99.759
4.5807 | -99.656
4.5844 | -99.5483
4.5881 | -99.4347
4.5919 | -99.3198
4.5956 | -99.208
4.5993 | -99.1036
4.6031 | -99.0195
4.6068 | -98.9692
4.6105 | -98.9559
4.6142 | -98.983
4.618 | -99.0509
4.6217 | -99.1553
4.6254 | -99.2954
4.6291 | -99.4575
4.6329 | -99.6197
4.6366 | -99.7857
4.6403 | -99.927
4.6441 | -100.0283
4.6478 | -100.0763
4.6515 | -100.0697
4.6552 | -100.0107
4.659 | -99.9193
4.6627 | -99.8201
4.6664 | -99.7302
4.6701 | -99.6676
4.6739 | -99.6374
4.6776 | -99.6241
4.6813 | -99.629
4.6851 | -99.6505
4.6888 | -99.6673
4.6925 | -99.6633
4.6962 | -99.6371
4.7 | -99.5987
4.7037 | -99.5533
4.7074 | -99.5046
4.7111 | -99.4648
4.7149 | -99.4322
4.7186 | -99.4112
4.7223 | -99.4023
4.7261 | -99.4101
4.7298 | -99.4208
4.7335 | -99.4222
4.7372 | -99.4153
4.741 | -99.3993
4.7447 | -99.3641
4.7484 | -99.3206
4.7521 | -99.2773
4.7559 | -99.2399
4.7596 | -99.2179
4.7633 | -99.2186
4.7671 | -99.229
4.7708 | -99.2332
4.7745 | -99.2431
4.7782 | -99.2609
4.782 | -99.2892
4.7857 | -99.325
4.7894 | -99.3625
4.7931 | -99.3998
4.7969 | -99.4298
4.8006 | -99.4711
4.8043 | -99.5112
4.8081 | -99.5414
4.8118 | -99.5573
4.8155 | -99.5663
4.8192 | -99.5533
4.823 | -99.5284
4.8267 | -99.5032
4.8304 | -99.4889
4.8341 | -99.4801
4.8379 | -99.5024
4.8416 | -99.5199
4.8453 | -99.5575
4.8491 | -99.6013
4.8528 | -99.6528
4.8565 | -99.7221
4.8602 | -99.8117
4.864 | -99.917
4.8677 | -100.0365
4.8714 | -100.159
4.8751 | -100.2644
4.8789 | -100.3549
4.8826 | -100.4155
4.8863 | -100.4505
4.89 | -100.4529
4.8938 | -100.4419
4.8975 | -100.4086
4.9012 | -100.3662
4.905 | -100.3149
4.9087 | -100.2665
4.9124 | -100.2128
4.9161 | -100.1643
4.9199 | -100.1194
4.9236 | -100.0804
4.9273 | -100.061
4.931 | -100.0582
4.9348 | -100.0746
4.9385 | -100.1076
4.9422 | -100.1534
4.946 | -100.2052
4.9497 | -100.2458
4.9534 | -100.2739
4.9571 | -100.2919
4.9609 | -100.2845
4.9646 | -100.2429
4.9683 | -100.192
4.972 | -100.1241
4.9758 | -100.0547
4.9795 | -99.9795
4.9832 | -99.9091
4.987 | -99.8441
4.9907 | -99.7902
4.9944 | -99.753
4.9981 | -99.7344
5.0019 | -99.7345
5.0056 | -99.7514
5.0093 | -99.7928
5.013 | -99.8445
5.0168 | -99.9039
5.0205 | -99.9633
5.0242 | -100.0148
5.028 | -100.0596
5.0317 | -100.0872
5.0354 | -100.1042
5.0391 | -100.1139
5.0429 | -100.1212
5.0466 | -100.113
5.0503 | -100.0993
5.054 | -100.0782
5.0578 | -100.0593
5.0615 | -100.0302
5.0652 | -99.9819
5.069 | -99.9248
5.0727 | -99.8531
5.0764 | -99.7798
5.0801 | -99.714
5.0839 | -99.6585
5.0876 | -99.608
5.0913 | -99.5718
5.095 | -99.5481
5.0988 | -99.5413
5.1025 | -99.5471
5.1062 | -99.5671
5.11 | -99.5986
5.1137 | -99.6362
5.1174 | -99.6851
5.1211 | -99.7335
5.1249 | -99.777
5.1286 | -99.8156
5.1323 | -99.8353
5.136 | -99.8412
5.1398 | -99.85
5.1435 | -99.8624
5.1472 | -99.8757
5.1509 | -99.9014
5.1547 | -99.9377
5.1584 | -99.9707
5.1621 | -100.0055
5.1659 | -100.0441
5.1696 | -100.0763
5.1733 | -100.1099
5.177 | -100.1351
5.1808 | -100.1579
5.1845 | -100.183
5.1882 | -100.2062
5.1919 | -100.2216
5.1957 | -100.2348
5.1994 | -100.2317
5.2031 | -100.2228
5.2069 | -100.2087
5.2106 | -100.1964
5.2143 | -100.1812
5.218 | -100.1613
5.2218 | -100.1519
5.2255 | -100.1461
5.2292 | -100.1415
5.2329 | -100.1437
5.2367 | -100.1551
5.2404 | -100.1591
5.2441 | -100.1628
5.2479 | -100.1734
5.2516 | -100.1856
5.2553 | -100.2047
5.259 | -100.2282
5.2628 | -100.2545
5.2665 | -100.2827
5.2702 | -100.3223
5.2739 | -100.3681
5.2777 | -100.4187
5.2814 | -100.4597
5.2851 | -100.4961
5.2889 | -100.5224
5.2926 | -100.5356
5.2963 | -100.5349
5.3 | -100.5208
5.3038 | -100.4963
5.3075 | -100.4662
5.3112 | -100.4263
5.3149 | -100.3846
5.3187 | -100.3548
5.3224 | -100.3299
5.3261 | -100.3109
5.3299 | -100.3026
5.3336 | -100.3077
5.3373 | -100.3247
5.341 | -100.3551
5.3448 | -100.3997
5.3485 | -100.4433
5.3522 | -100.4931
5.3559 | -100.5569
5.3597 | -100.6291
5.3634 | -100.6924
5.3671 | -100.7392
5.3709 | -100.7622
5.3746 | -100.73
5.3783 | -100.7089
5.382 | -100.6727
5.3858 | -100.6282
5.3895 | -100.5786
5.3932 | -100.5314
5.3969 | -100.4832
5.4007 | -100.4382
5.4044 | -100.4012
5.4081 | -100.3661
5.4119 | -100.3359
5.4156 | -100.3191
5.4193 | -100.3132
5.423 | -100.3099
5.4268 | -100.3132
5.4305 | -100.3293
5.4342 | -100.3472
5.4379 | -100.3652
5.4417 | -100.3916
5.4454 | -100.4158
5.4491 | -100.4407
5.4529 | -100.4652
5.4566 | -100.4852
5.4603 | -100.5055
5.464 | -100.5249
5.4678 | -100.5226
5.4715 | -100.5118
5.4752 | -100.4978
5.4789 | -100.4784
5.4827 | -100.455
5.4864 | -100.4187
5.4901 | -100.3795
5.4939 | -100.346
5.4976 | -100.3161
5.5013 | -100.2866
5.505 | -100.274
5.5088 | -100.2804
5.5125 | -100.295
5.5162 | -100.3158
5.5199 | -100.3511
5.5237 | -100.394
5.5274 | -100.4295
5.5311 | -100.4746
5.5348 | -100.5135
5.5386 | -100.5525
5.5423 | -100.5928
5.546 | -100.6278
5.5498 | -100.6638
5.5535 | -100.7056
5.5572 | -100.7348
5.5609 | -100.7482
5.5647 | -100.758
5.5684 | -100.761
5.5721 | -100.7623
5.5758 | -100.7556
5.5796 | -100.7526
5.5833 | -100.7459
5.587 | -100.7341
5.5908 | -100.7148
5.5945 | -100.699
5.5982 | -100.6898
5.6019 | -100.6812
5.6057 | -100.6732
5.6094 | -100.6733
5.6131 | -100.686
5.6168 | -100.7082
5.6206 | -100.7333
5.6243 | -100.7538
5.628 | -100.7753
5.6318 | -100.7895
5.6355 | -100.7901
5.6392 | -100.7779
5.6429 | -100.7451
5.6467 | -100.6995
5.6504 | -100.6447
5.6541 | -100.5842
5.6578 | -100.5193
5.6616 | -100.4615
5.6653 | -100.4139
5.669 | -100.3844
5.6728 | -100.3731
5.6765 | -100.3785
5.6802 | -100.3971
5.6839 | -100.4337
5.6877 | -100.4805
5.6914 | -100.5367
5.6951 | -100.6057
5.6988 | -100.6813
5.7026 | -100.7614
5.7063 | -100.8564
5.71 | -100.9502
5.7138 | -101.0398
5.7175 | -101.1386
5.7212 | -101.2326
5.7249 | -101.3144
5.7287 | -101.3962
5.7324 | -101.4678
5.7361 | -101.5218
5.7398 | -101.5616
5.7436 | -101.5691
5.7473 | -101.5498
5.751 | -101.5008
5.7548 | -101.4264
5.7585 | -101.3379
5.7622 | -101.239
5.7659 | -101.1304
5.7697 | -101.0225
5.7734 | -100.9331
5.7771 | -100.8577
5.7808 | -100.7947
5.7846 | -100.749
5.7883 | -100.7232
5.792 | -100.7087
5.7958 | -100.7061
5.7995 | -100.7212
5.8032 | -100.7463
5.8069 | -100.7772
5.8107 | -100.8097
5.8144 | -100.8377
5.8181 | -100.8599
5.8218 | -100.875
5.8256 | -100.8761
5.8293 | -100.8739
5.833 | -100.8559
5.8368 | -100.8321
5.8405 | -100.81
5.8442 | -100.792
5.8479 | -100.7812
5.8517 | -100.7876
5.8554 | -100.8016
5.8591 | -100.8154
5.8628 | -100.8351
5.8666 | -100.8519
5.8703 | -100.8651
5.874 | -100.8698
5.8777 | -100.8668
5.8815 | -100.8577
5.8852 | -100.8337
5.8889 | -100.8018
5.8927 | -100.7771
5.8964 | -100.7576
5.9001 | -100.7506
5.9038 | -100.7529
5.9076 | -100.7741
5.9113 | -100.8267
5.915 | -100.8573
5.9187 | -100.8895
5.9225 | -100.9212
5.9262 | -100.9493
5.9299 | -100.9697
5.9337 | -100.9672
5.9374 | -100.9594
5.9411 | -100.9433
5.9448 | -100.9239
5.9486 | -100.9
5.9523 | -100.8705
5.956 | -100.8496
5.9597 | -100.8307
5.9635 | -100.8128
5.9672 | -100.8005
5.9709 | -100.7866
5.9747 | -100.7741
5.9784 | -100.7683
5.9821 | -100.764
5.9858 | -100.7657
5.9896 | -100.7693
5.9933 | -100.7849
5.997 | -100.797
6.0007 | -100.8179
6.0045 | -100.8428
6.0082 | -100.8704
6.0119 | -100.8976
6.0157 | -100.9221
6.0194 | -100.9422
6.0231 | -100.9644
6.0268 | -100.9889
6.0306 | -101.0087
6.0343 | -101.0308
6.038 | -101.056
6.0417 | -101.0862
6.0455 | -101.1134
6.0492 | -101.1418
6.0529 | -101.1757
6.0567 | -101.2076
6.0604 | -101.2329
6.0641 | -101.2592
6.0678 | -101.2785
6.0716 | -101.2908
6.0753 | -101.3046
6.079 | -101.3183
6.0827 | -101.331
6.0865 | -101.3396
6.0902 | -101.3562
6.0939 | -101.3734
6.0977 | -101.395
6.1014 | -101.4174
6.1051 | -101.4319
6.1088 | -101.448
6.1126 | -101.4602
6.1163 | -101.4582
6.12 | -101.4514
6.1237 | -101.4325
6.1275 | -101.4005
6.1312 | -101.3639
6.1349 | -101.3207
6.1387 | -101.2727
6.1424 | -101.2081
6.1461 | -101.1388
6.1498 | -101.0668
6.1536 | -101.0024
6.1573 | -100.9385
6.161 | -100.8856
6.1647 | -100.8387
6.1685 | -100.7999
6.1722 | -100.7773
6.1759 | -100.775
6.1796 | -100.7879
6.1834 | -100.8122
6.1871 | -100.8436
6.1908 | -100.8784
6.1946 | -100.9151
6.1983 | -100.9487
6.202 | -100.9741
6.2057 | -100.9998
6.2095 | -101.0275
6.2132 | -101.0447
6.2169 | -101.0556
6.2206 | -101.0648
6.2244 | -101.0625
6.2281 | -101.0598
6.2318 | -101.0618
6.2356 | -101.0677
6.2393 | -101.0788
6.243 | -101.1036
6.2467 | -101.1377
6.2505 | -101.1856
6.2542 | -101.2443
6.2579 | -101.3181
6.2616 | -101.3958
6.2654 | -101.4696
6.2691 | -101.5411
6.2728 | -101.6033
6.2766 | -101.6505
6.2803 | -101.6886
6.284 | -101.7177
6.2877 | -101.7402
6.2915 | -101.7586
6.2952 | -101.7717
6.2989 | -101.7714
6.3026 | -101.7585
6.3064 | -101.7467
6.3101 | -101.7345
6.3138 | -101.7156
6.3176 | -101.6922
6.3213 | -101.6676
6.325 | -101.6378
6.3287 | -101.6044
6.3325 | -101.5653
6.3362 | -101.5305
6.3399 | -101.4954
6.3436 | -101.4672
6.3474 | -101.4515
6.3511 | -101.4461
6.3548 | -101.4542
6.3586 | -101.4754
6.3623 | -101.5054
6.366 | -101.5455
6.3697 | -101.5937
6.3735 | -101.6364
6.3772 | -101.6771
6.3809 | -101.7138
6.3846 | -101.7464
6.3884 | -101.7793
6.3921 | -101.8193
6.3958 | -101.8647
6.3996 | -101.9089
6.4033 | -101.95
6.407 | -101.9889
6.4107 | -102.0255
6.4145 | -102.0513
6.4182 | -102.0707
6.4219 | -102.0872
6.4256 | -102.0948
6.4294 | -102.0892
6.4331 | -102.0836
6.4368 | -102.0686
6.4406 | -102.0428
6.4443 | -102.0198
6.448 | -101.9998
6.4517 | -101.9684
6.4555 | -101.9363
6.4592 | -101.904
6.4629 | -101.8727
6.4666 | -101.8392
6.4704 | -101.81
6.4741 | -101.7857
6.4778 | -101.7627
6.4816 | -101.7435
6.4853 | -101.7262
6.489 | -101.7105
6.4927 | -101.6969
6.4965 | -101.6853
6.5002 | -101.6797
6.5039 | -101.6806
6.5076 | -101.6809
6.5114 | -101.6787
6.5151 | -101.6734
6.5188 | -101.6642
6.5225 | -101.657
6.5263 | -101.6482
6.53 | -101.6426
6.5337 | -101.6394
6.5375 | -101.6402
6.5412 | -101.6441
6.5449 | -101.652
6.5486 | -101.6693
6.5524 | -101.6924
6.5561 | -101.7162
6.5598 | -101.7337
6.5635 | -101.7479
6.5673 | -101.761
6.571 | -101.771
6.5747 | -101.774
6.5785 | -101.7755
6.5822 | -101.7738
6.5859 | -101.7676
6.5896 | -101.7617
6.5934 | -101.7561
6.5971 | -101.7476
6.6008 | -101.7346
6.6045 | -101.7206
6.6083 | -101.7125
6.612 | -101.709
6.6157 | -101.7111
6.6195 | -101.7206
6.6232 | -101.7346
6.6269 | -101.7494
6.6306 | -101.7639
6.6344 | -101.784
6.6381 | -101.7997
6.6418 | -101.8096
6.6455 | -101.817
6.6493 | -101.8233
6.653 | -101.8332
6.6567 | -101.8383
6.6605 | -101.8448
6.6642 | -101.8501
6.6679 | -101.8531
6.6716 | -101.8538
6.6754 | -101.856
6.6791 | -101.8592
6.6828 | -101.865
6.6865 | -101.8754
6.6903 | -101.894
6.694 | -101.9162
6.6977 | -101.9444
6.7015 | -101.9756
6.7052 | -102.0009
6.7089 | -102.0234
6.7126 | -102.038
6.7164 | -102.0401
6.7201 | -102.0339
6.7238 | -102.0123
6.7275 | -101.9787
6.7313 | -101.9356
6.735 | -101.8826
6.7387 | -101.824
6.7425 | -101.7661
6.7462 | -101.7095
6.7499 | -101.6535
6.7536 | -101.6063
6.7574 | -101.5733
6.7611 | -101.5476
6.7648 | -101.5295
6.7685 | -101.5189
6.7723 | -101.5124
6.776 | -101.5073
6.7797 | -101.4995
6.7835 | -101.4891
6.7872 | -101.4817
6.7909 | -101.4726
6.7946 | -101.4558
6.7984 | -101.4326
6.8021 | -101.4049
6.8058 | -101.3834
6.8095 | -101.3612
6.8133 | -101.3353
6.817 | -101.3132
6.8207 | -101.2924
6.8244 | -101.2776
6.8282 | -101.2686
6.8319 | -101.2712
6.8356 | -101.2868
6.8394 | -101.315
6.8431 | -101.3522
6.8468 | -101.3953
6.8505 | -101.4388
6.8543 | -101.4796
6.858 | -101.5196
6.8617 | -101.56
6.8654 | -101.5942
6.8692 | -101.6213
6.8729 | -101.6406
6.8766 | -101.6507
6.8804 | -101.6629
6.8841 | -101.6758
6.8878 | -101.6831
6.8915 | -101.6885
6.8953 | -101.7097
6.899 | -101.7355
6.9027 | -101.768
6.9064 | -101.7991
6.9102 | -101.8312
6.9139 | -101.8777
6.9176 | -101.9222
6.9214 | -101.9691
6.9251 | -102.0182
6.9288 | -102.0679
6.9325 | -102.105
6.9363 | -102.1381
6.94 | -102.1674
6.9437 | -102.1873
6.9474 | -102.203
6.9512 | -102.2131
6.9549 | -102.2187
6.9586 | -102.2191
6.9624 | -102.2135
6.9661 | -102.2072
6.9698 | -102.1983
6.9735 | -102.1862
6.9773 | -102.1757
6.981 | -102.154
6.9847 | -102.1305
6.9884 | -102.1172
6.9922 | -102.0967
6.9959 | -102.0859
6.9996 | -102.0739
7.0034 | -102.0659
7.0071 | -102.0607
7.0108 | -102.0596
7.0145 | -102.0616
7.0183 | -102.0691
7.022 | -102.0852
7.0257 | -102.1003
7.0294 | -102.1174
7.0332 | -102.1338
7.0369 | -102.1551
7.0406 | -102.1783
7.0444 | -102.2027
7.0481 | -102.2297
7.0518 | -102.2557
7.0555 | -102.2825
7.0593 | -102.3061
7.063 | -102.3247
7.0667 | -102.3444
7.0704 | -102.3605
7.0742 | -102.3732
7.0779 | -102.384
7.0816 | -102.3899
7.0854 | -102.393
7.0891 | -102.3909
7.0928 | -102.3882
7.0965 | -102.3964
7.1003 | -102.407
7.104 | -102.4143
7.1077 | -102.4265
7.1114 | -102.4416
7.1152 | -102.459
7.1189 | -102.4766
7.1226 | -102.4884
7.1264 | -102.4891
7.1301 | -102.4831
7.1338 | -102.4657
7.1375 | -102.4429
7.1413 | -102.4172
7.145 | -102.3833
7.1487 | -102.3454
7.1524 | -102.3044
7.1562 | -102.2619
7.1599 | -102.2262
7.1636 | -102.1908
7.1673 | -102.1577
7.1711 | -102.131
7.1748 | -102.1128
7.1785 | -102.1076
7.1823 | -102.1109
7.186 | -102.1241
7.1897 | -102.1442
7.1934 | -102.1785
7.1972 | -102.2167
7.2009 | -102.2589
7.2046 | -102.3073
7.2083 | -102.3578
7.2121 | -102.4102
7.2158 | -102.4661
7.2195 | -102.5148
7.2233 | -102.5553
7.227 | -102.5939
7.2307 | -102.6226
7.2344 | -102.6382
7.2382 | -102.6466
7.2419 | -102.6442
7.2456 | -102.6272
7.2493 | -102.5996
7.2531 | -102.5665
7.2568 | -102.525
7.2605 | -102.4825
7.2643 | -102.4445
7.268 | -102.4008
7.2717 | -102.3669
7.2754 | -102.3357
7.2792 | -102.3096
7.2829 | -102.2829
7.2866 | -102.259
7.2903 | -102.2467
7.2941 | -102.2394
7.2978 | -102.2372
7.3015 | -102.2427
7.3053 | -102.2574
7.309 | -102.2734
7.3127 | -102.2941
7.3164 | -102.3243
7.3202 | -102.3583
7.3239 | -102.3942
7.3276 | -102.4382
7.3313 | -102.4813
7.3351 | -102.5239
7.3388 | -102.5707
7.3425 | -102.6217
7.3463 | -102.6687
7.35 | -102.7167
7.3537 | -102.7581
7.3574 | -102.7993
7.3612 | -102.8395
7.3649 | -102.8793
7.3686 | -102.9118
7.3723 | -102.9348
7.3761 | -102.9563
7.3798 | -102.9645
7.3835 | -102.9657
7.3873 | -102.9588
7.391 | -102.9373
7.3947 | -102.8999
7.3984 | -102.8546
7.4022 | -102.8084
7.4059 | -102.7551
7.4096 | -102.69
7.4133 | -102.6314
7.4171 | -102.5821
7.4208 | -102.5374
7.4245 | -102.4965
7.4283 | -102.4593
7.432 | -102.4308
7.4357 | -102.4178
7.4394 | -102.415
7.4432 | -102.4205
7.4469 | -102.4373
7.4506 | -102.4579
7.4543 | -102.4846
7.4581 | -102.5151
7.4618 | -102.5498
7.4655 | -102.5753
7.4693 | -102.5961
7.473 | -102.615
7.4767 | -102.6238
7.4804 | -102.6228
7.4842 | -102.613
7.4879 | -102.5977
7.4916 | -102.5795
7.4953 | -102.5566
7.4991 | -102.5392
7.5028 | -102.5281
7.5065 | -102.5146
7.5103 | -102.5033
7.514 | -102.4988
7.5177 | -102.4958
7.5214 | -102.4904
7.5252 | -102.5078
7.5289 | -102.5151
7.5326 | -102.5259
7.5363 | -102.5447
7.5401 | -102.5725
7.5438 | -102.6029
7.5475 | -102.6283
7.5512 | -102.661
7.555 | -102.6894
7.5587 | -102.7136
7.5624 | -102.7396
7.5662 | -102.7641
7.5699 | -102.7829
7.5736 | -102.7969
7.5773 | -102.8068
7.5811 | -102.8107
7.5848 | -102.8039
7.5885 | -102.7977
7.5922 | -102.7895
7.596 | -102.7755
7.5997 | -102.7596
7.6034 | -102.747
7.6072 | -102.7389
7.6109 | -102.7342
7.6146 | -102.7271
7.6183 | -102.7256
7.6221 | -102.7348
7.6258 | -102.7554
7.6295 | -102.7834
7.6332 | -102.8101
7.637 | -102.8336
7.6407 | -102.8552
7.6444 | -102.8787
7.6482 | -102.9009
7.6519 | -102.9179
7.6556 | -102.9375
7.6593 | -102.9519
7.6631 | -102.9649
7.6668 | -102.9678
7.6705 | -102.9673
7.6742 | -102.9625
7.678 | -102.9513
7.6817 | -102.9402
7.6854 | -102.9282
7.6892 | -102.9171
7.6929 | -102.9137
7.6966 | -102.9105
7.7003 | -102.9111
7.7041 | -102.9212
7.7078 | -102.9351
7.7115 | -102.9576
7.7152 | -102.9871
7.719 | -103.0166
7.7227 | -103.0441
7.7264 | -103.0689
7.7302 | -103.0954
7.7339 | -103.1173
7.7376 | -103.1299
7.7413 | -103.1376
7.7451 | -103.1457
7.7488 | -103.1509
7.7525 | -103.1536
7.7562 | -103.1505
7.76 | -103.141
7.7637 | -103.1256
7.7674 | -103.104
7.7712 | -103.0734
7.7749 | -103.039
7.7786 | -103.0034
7.7823 | -102.9646
7.7861 | -102.9224
7.7898 | -102.874
7.7935 | -102.8268
7.7972 | -102.7798
7.801 | -102.7313
7.8047 | -102.6844
7.8084 | -102.6419
7.8122 | -102.6119
7.8159 | -102.5891
7.8196 | -102.5727
7.8233 | -102.566
7.8271 | -102.5655
7.8308 | -102.5766
7.8345 | -102.5963
7.8382 | -102.6268
7.842 | -102.6605
7.8457 | -102.6971
7.8494 | -102.7315
7.8531 | -102.7707
7.8569 | -102.8033
7.8606 | -102.8283
7.8643 | -102.8505
7.8681 | -102.8685
7.8718 | -102.8842
7.8755 | -102.903
7.8792 | -102.9203
7.883 | -102.9334
7.8867 | -102.949
7.8904 | -102.9611
7.8941 | -102.973
7.8979 | -102.9839
7.9016 | -102.9846
7.9053 | -102.987
7.9091 | -102.986
7.9128 | -102.9841
7.9165 | -102.9895
7.9202 | -102.991
7.924 | -102.995
7.9277 | -103.001
7.9314 | -103.0101
7.9351 | -103.0226
7.9389 | -103.0447
7.9426 | -103.0734
7.9463 | -103.1007
7.9501 | -103.1322
7.9538 | -103.1705
7.9575 | -103.2175
7.9612 | -103.2647
7.965 | -103.3156
7.9687 | -103.3533
7.9724 | -103.39
7.9761 | -103.4205
7.9799 | -103.4443
7.9836 | -103.4544
7.9873 | -103.452
7.9911 | -103.4427
7.9948 | -103.4238
7.9985 | -103.401
8.0022 | -103.3687
8.006 | -103.335
8.0097 | -103.2953
8.0134 | -103.2569
8.0171 | -103.2213
8.0209 | -103.1789
8.0246 | -103.1331
8.0283 | -103.0876
8.0321 | -103.0397
8.0358 | -102.9972
8.0395 | -102.9538
8.0432 | -102.9103
8.047 | -102.8727
8.0507 | -102.8378
8.0544 | -102.7987
8.0581 | -102.762
8.0619 | -102.7354
8.0656 | -102.7198
8.0693 | -102.7002
8.0731 | -102.6829
8.0768 | -102.6674
8.0805 | -102.6606
8.0842 | -102.6561
8.088 | -102.661
8.0917 | -102.6737
8.0954 | -102.6907
8.0991 | -102.7067
8.1029 | -102.7283
8.1066 | -102.7557
8.1103 | -102.7822
8.1141 | -102.8149
8.1178 | -102.8483
8.1215 | -102.8815
8.1252 | -102.914
8.129 | -102.9456
8.1327 | -102.9747
8.1364 | -102.9971
8.1401 | -103.0146
8.1439 | -103.0297
8.1476 | -103.0403
8.1513 | -103.046
8.155 | -103.0504
8.1588 | -103.0559
8.1625 | -103.0646
8.1662 | -103.0727
8.17 | -103.0798
8.1737 | -103.0939
8.1774 | -103.1085
8.1811 | -103.1211
8.1849 | -103.1378
8.1886 | -103.1644
8.1923 | -103.1891
8.196 | -103.2214
8.1998 | -103.2607
8.2035 | -103.2951
8.2072 | -103.3307
8.211 | -103.3648
8.2147 | -103.3895
8.2184 | -103.416
8.2221 | -103.4375
8.2259 | -103.4503
8.2296 | -103.455
8.2333 | -103.4483
8.237 | -103.4369
8.2408 | -103.4215
8.2445 | -103.402
8.2482 | -103.3838
8.252 | -103.3626
8.2557 | -103.339
8.2594 | -103.3232
8.2631 | -103.3103
8.2669 | -103.3027
8.2706 | -103.3051
8.2743 | -103.318
8.278 | -103.335
8.2818 | -103.3588
8.2855 | -103.3844
8.2892 | -103.4172
8.293 | -103.4608
8.2967 | -103.5061
8.3004 | -103.5493
8.3041 | -103.5862
8.3079 | -103.6221
8.3116 | -103.6527
8.3153 | -103.6775
8.319 | -103.6937
8.3228 | -103.7015
8.3265 | -103.7099
8.3302 | -103.7116
8.334 | -103.7033
8.3377 | -103.6957
8.3414 | -103.6835
8.3451 | -103.6707
8.3489 | -103.6593
8.3526 | -103.6458
8.3563 | -103.641
8.36 | -103.6371
8.3638 | -103.632
8.3675 | -103.6284
8.3712 | -103.6289
8.375 | -103.6375
8.3787 | -103.6488
8.3824 | -103.6578
8.3861 | -103.6624
8.3899 | -103.6603
8.3936 | -103.6535
8.3973 | -103.6432
8.401 | -103.6304
8.4048 | -103.6166
8.4085 | -103.6067
8.4122 | -103.59
8.416 | -103.5654
8.4197 | -103.5419
8.4234 | -103.5236
8.4271 | -103.5081
8.4309 | -103.4867
8.4346 | -103.4642
8.4383 | -103.4498
8.442 | -103.4326
8.4458 | -103.4175
8.4495 | -103.4017
8.4532 | -103.3873
8.457 | -103.3795
8.4607 | -103.3674
8.4644 | -103.3516
8.4681 | -103.3383
8.4719 | -103.3294
8.4756 | -103.3181
8.4793 | -103.3099
8.483 | -103.3029
8.4868 | -103.3019
8.4905 | -103.3082
8.4942 | -103.3123
8.4979 | -103.318
8.5017 | -103.3286
8.5054 | -103.3359
8.5091 | -103.3492
8.5129 | -103.3599
8.5166 | -103.3683
8.5203 | -103.3753
8.524 | -103.3782
8.5278 | -103.372
8.5315 | -103.3635
8.5352 | -103.3555
8.5389 | -103.3479
8.5427 | -103.3421
8.5464 | -103.3358
8.5501 | -103.3254
8.5539 | -103.3179
8.5576 | -103.3133
8.5613 | -103.306
8.565 | -103.3017
8.5688 | -103.3003
8.5725 | -103.3011
8.5762 | -103.3015
8.5799 | -103.2984
8.5837 | -103.2982
8.5874 | -103.303
8.5911 | -103.3091
8.5949 | -103.3168
8.5986 | -103.3285
8.6023 | -103.3409
8.606 | -103.3501
8.6098 | -103.3662
8.6135 | -103.3836
8.6172 | -103.3997
8.6209 | -103.4228
8.6247 | -103.4455
8.6284 | -103.4631
8.6321 | -103.4819
8.6359 | -103.4952
8.6396 | -103.5046
8.6433 | -103.5135
8.647 | -103.518
8.6508 | -103.5143
8.6545 | -103.5057
8.6582 | -103.4901
8.6619 | -103.469
8.6657 | -103.4439
8.6694 | -103.4119
8.6731 | -103.379
8.6769 | -103.3469
8.6806 | -103.3089
8.6843 | -103.2741
8.688 | -103.2397
8.6918 | -103.2048
8.6955 | -103.1676
8.6992 | -103.1338
8.7029 | -103.1035
8.7067 | -103.0781
8.7104 | -103.0545
8.7141 | -103.0339
8.7179 | -103.0186
8.7216 | -103.0099
8.7253 | -103.0108
8.729 | -103.0151
8.7328 | -103.0226
8.7365 | -103.0276
8.7402 | -103.0304
8.7439 | -103.036
8.7477 | -103.0417
8.7514 | -103.049
8.7551 | -103.061
8.7589 | -103.0753
8.7626 | -103.0913
8.7663 | -103.1075
8.77 | -103.1264
8.7738 | -103.1444
8.7775 | -103.1663
8.7812 | -103.1909
8.7849 | -103.2146
8.7887 | -103.2379
8.7924 | -103.2641
8.7961 | -103.2848
8.7999 | -103.3082
8.8036 | -103.3307
8.8073 | -103.3507
8.811 | -103.3738
8.8148 | -103.3885
8.8185 | -103.4043
8.8222 | -103.4212
8.8259 | -103.4357
8.8297 | -103.4499
8.8334 | -103.4704
8.8371 | -103.4916
8.8408 | -103.5191
8.8446 | -103.5491
8.8483 | -103.5815
8.852 | -103.6161
8.8558 | -103.6576
8.8595 | -103.6988
8.8632 | -103.7401
8.8669 | -103.7868
8.8707 | -103.8337
8.8744 | -103.8729
8.8781 | -103.9059
8.8818 | -103.9292
8.8856 | -103.9434
8.8893 | -103.9579
8.893 | -103.962
8.8968 | -103.957
8.9005 | -103.9472
8.9042 | -103.9339
8.9079 | -103.9172
8.9117 | -103.9004
8.9154 | -103.8857
8.9191 | -103.864
8.9228 | -103.8374
8.9266 | -103.8156
8.9303 | -103.7885
8.934 | -103.7573
8.9378 | -103.7288
8.9415 | -103.7096
8.9452 | -103.6979
8.9489 | -103.6866
8.9527 | -103.6758
8.9564 | -103.6761
8.9601 | -103.6751
8.9638 | -103.6689
8.9676 | -103.6651
8.9713 | -103.6584
8.975 | -103.6543
8.9788 | -103.6534
8.9825 | -103.649
8.9862 | -103.6379
8.9899 | -103.6328
8.9937 | -103.6315
8.9974 | -103.6307
9.0011 | -103.6322
9.0048 | -103.6428
9.0086 | -103.658
9.0123 | -103.6744
9.016 | -103.6944
9.0198 | -103.7123
9.0235 | -103.734
9.0272 | -103.7578
9.0309 | -103.7833
9.0347 | -103.8062
9.0384 | -103.8233
9.0421 | -103.8399
9.0458 | -103.8483
9.0496 | -103.8593
9.0533 | -103.8676
9.057 | -103.8718
9.0608 | -103.876
9.0645 | -103.8759
9.0682 | -103.8706
9.0719 | -103.8505
9.0757 | -103.8282
9.0794 | -103.8081
9.0831 | -103.7892
9.0868 | -103.7631
9.0906 | -103.7389
9.0943 | -103.7221
9.098 | -103.7072
9.1018 | -103.694
9.1055 | -103.6823
9.1092 | -103.6685
9.1129 | -103.6637
9.1167 | -103.6656
9.1204 | -103.6663
9.1241 | -103.6663
9.1278 | -103.6689
9.1316 | -103.6735
9.1353 | -103.6805
9.139 | -103.7014
9.1428 | -103.7081
9.1465 | -103.7182
9.1502 | -103.7309
9.1539 | -103.7516
9.1577 | -103.7748
9.1614 | -103.7958
9.1651 | -103.8119
9.1688 | -103.826
9.1726 | -103.8406
9.1763 | -103.8529
9.18 | -103.8618
9.1837 | -103.8592
9.1875 | -103.8541
9.1912 | -103.8453
9.1949 | -103.8315
9.1987 | -103.8186
9.2024 | -103.8047
9.2061 | -103.7875
9.2098 | -103.7726
9.2136 | -103.7579
9.2173 | -103.7456
9.221 | -103.7357
9.2247 | -103.7288
9.2285 | -103.7243
9.2322 | -103.7229
9.2359 | -103.7277
9.2397 | -103.7341
9.2434 | -103.7477
9.2471 | -103.7624
9.2508 | -103.7843
9.2546 | -103.8094
9.2583 | -103.8327
9.262 | -103.8516
9.2657 | -103.879
9.2695 | -103.9084
9.2732 | -103.9391
9.2769 | -103.969
9.2807 | -103.9965
9.2844 | -104.0131
9.2881 | -104.0282
9.2918 | -104.0333
9.2956 | -104.0344
9.2993 | -104.0328
9.303 | -104.0266
9.3067 | -104.0159
9.3105 | -103.9997
9.3142 | -103.9781
9.3179 | -103.9562
9.3217 | -103.932
9.3254 | -103.9021
9.3291 | -103.8753
9.3328 | -103.8496
9.3366 | -103.8192
9.3403 | -103.7874
9.344 | -103.7618
9.3477 | -103.7307
9.3515 | -103.7043
9.3552 | -103.6834
9.3589 | -103.6652
9.3627 | -103.6465
9.3664 | -103.63
9.3701 | -103.6175
9.3738 | -103.607
9.3776 | -103.5982
9.3813 | -103.5892
9.385 | -103.5829
9.3887 | -103.5788
9.3925 | -103.5795
9.3962 | -103.5839
9.3999 | -103.5941
9.4037 | -103.5973
9.4074 | -103.5992
9.4111 | -103.6024
9.4148 | -103.611
9.4186 | -103.6268
9.4223 | -103.6455
9.426 | -103.656
9.4297 | -103.6666
9.4335 | -103.6795
9.4372 | -103.6947
9.4409 | -103.7134
9.4447 | -103.7322
9.4484 | -103.7495
9.4521 | -103.7712
9.4558 | -103.792
9.4596 | -103.8184
9.4633 | -103.8439
9.467 | -103.8648
9.4707 | -103.8827
9.4745 | -103.9074
9.4782 | -103.9317
9.4819 | -103.9545
9.4856 | -103.9803
9.4894 | -103.999
9.4931 | -104.0146
9.4968 | -104.0326
9.5006 | -104.0544
9.5043 | -104.0737
9.508 | -104.0939
9.5117 | -104.1195
9.5155 | -104.1466
9.5192 | -104.1749
9.5229 | -104.2044
9.5266 | -104.2325
9.5304 | -104.2591
9.5341 | -104.2879
9.5378 | -104.3188
9.5416 | -104.3521
9.5453 | -104.3838
9.549 | -104.4105
9.5527 | -104.4331
9.5565 | -104.4506
9.5602 | -104.4638
9.5639 | -104.471
9.5676 | -104.4689
9.5714 | -104.46
9.5751 | -104.439
9.5788 | -104.4116
9.5826 | -104.3791
9.5863 | -104.3415
9.59 | -104.2939
9.5937 | -104.2416
9.5975 | -104.1842
9.6012 | -104.1318
9.6049 | -104.0817
9.6086 | -104.0349
9.6124 | -103.9904
9.6161 | -103.9518
9.6198 | -103.9192
9.6236 | -103.8967
9.6273 | -103.8861
9.631 | -103.8857
9.6347 | -103.8942
9.6385 | -103.906
9.6422 | -103.9275
9.6459 | -103.9545
9.6496 | -103.9827
9.6534 | -104.0168
9.6571 | -104.0528
9.6608 | -104.0849
9.6646 | -104.1186
9.6683 | -104.1491
9.672 | -104.1739
9.6757 | -104.1691
9.6795 | -104.1798
9.6832 | -104.1868
9.6869 | -104.1902
9.6906 | -104.1905
9.6944 | -104.1842
9.6981 | -104.174
9.7018 | -104.1644
9.7056 | -104.1561
9.7093 | -104.1494
9.713 | -104.1403
9.7167 | -104.1324
9.7205 | -104.1253
9.7242 | -104.1175
9.7279 | -104.1109
9.7316 | -104.1059
9.7354 | -104.1065
9.7391 | -104.1045
9.7428 | -104.1008
9.7466 | -104.1029
9.7503 | -104.111
9.754 | -104.1155
9.7577 | -104.1221
9.7615 | -104.127
9.7652 | -104.1291
9.7689 | -104.1321
9.7726 | -104.136
9.7764 | -104.1365
9.7801 | -104.1383
9.7838 | -104.1388
9.7876 | -104.1378
9.7913 | -104.1357
9.795 | -104.1309
9.7987 | -104.1218
9.8025 | -104.1159
9.8062 | -104.1078
9.8099 | -104.0999
9.8136 | -104.0915
9.8174 | -104.0867
9.8211 | -104.0802
9.8248 | -104.0768
9.8285 | -104.0752
9.8323 | -104.0768
9.836 | -104.0807
9.8397 | -104.0907
9.8435 | -104.102
9.8472 | -104.1128
9.8509 | -104.1231
9.8546 | -104.1353
9.8584 | -104.1482
9.8621 | -104.1604
9.8658 | -104.1707
9.8695 | -104.1882
9.8733 | -104.2076
9.877 | -104.2221
9.8807 | -104.2276
9.8845 | -104.24
9.8882 | -104.2493
9.8919 | -104.2512
9.8956 | -104.2493
9.8994 | -104.2491
9.9031 | -104.2431
9.9068 | -104.236
9.9105 | -104.2287
9.9143 | -104.2195
9.918 | -104.2058
9.9217 | -104.1953
9.9255 | -104.1855
9.9292 | -104.1762
9.9329 | -104.1706
9.9366 | -104.1682
9.9404 | -104.1706
9.9441 | -104.1811
9.9478 | -104.1951
9.9515 | -104.2138
9.9553 | -104.236
9.959 | -104.2653
9.9627 | -104.2973
9.9665 | -104.3354
9.9702 | -104.3732
9.9739 | -104.4104
9.9776 | -104.449
9.9814 | -104.4874
9.9851 | -104.5251
9.9888 | -104.5633
9.9925 | -104.6026
9.9963 | -104.6371
10 | -104.6659
10.0037 | -104.6877
10.0075 | -104.7076
10.0112 | -104.7194
10.0149 | -104.7182
10.0186 | -104.712
10.0224 | -104.7047
10.0261 | -104.6941
10.0298 | -104.6683
10.0335 | -104.6392
10.0373 | -104.6095
10.041 | -104.5789
10.0447 | -104.5454
10.0485 | -104.5126
10.0522 | -104.4793
10.0559 | -104.4438
10.0596 | -104.4092
10.0634 | -104.3832
10.0671 | -104.356
10.0708 | -104.3328
10.0745 | -104.3142
10.0783 | -104.2999
10.082 | -104.2853
10.0857 | -104.2761
10.0895 | -104.2715
10.0932 | -104.269
10.0969 | -104.2693
10.1006 | -104.2742
10.1044 | -104.2786
10.1081 | -104.2898
10.1118 | -104.3034
10.1155 | -104.3239
10.1193 | -104.3503
10.123 | -104.381
10.1267 | -104.4126
10.1305 | -104.4411
10.1342 | -104.4733
10.1379 | -104.504
10.1416 | -104.532
10.1454 | -104.557
10.1491 | -104.5845
10.1528 | -104.6095
10.1565 | -104.6344
10.1603 | -104.6538
10.164 | -104.667
10.1677 | -104.6759
10.1715 | -104.6843
10.1752 | -104.6907
10.1789 | -104.6912
10.1826 | -104.6868
10.1864 | -104.6776
10.1901 | -104.6692
10.1938 | -104.6605
10.1975 | -104.6428
10.2013 | -104.6246
10.205 | -104.6087
10.2087 | -104.5894
10.2124 | -104.5738
10.2162 | -104.5675
10.2199 | -104.5558
10.2236 | -104.5408
10.2274 | -104.5274
10.2311 | -104.5201
10.2348 | -104.5094
10.2385 | -104.4983
10.2423 | -104.4917
10.246 | -104.4852
10.2497 | -104.4764
10.2534 | -104.4688
10.2572 | -104.458
10.2609 | -104.4455
10.2646 | -104.4319
10.2684 | -104.4172
10.2721 | -104.4019
10.2758 | -104.386
10.2795 | -104.3738
10.2833 | -104.3563
10.287 | -104.3374
10.2907 | -104.3202
10.2944 | -104.3038
10.2982 | -104.2852
10.3019 | -104.2669
10.3056 | -104.2496
10.3094 | -104.2369
10.3131 | -104.2236
10.3168 | -104.2099
10.3205 | -104.1986
10.3243 | -104.1885
10.328 | -104.1799
10.3317 | -104.1707
10.3354 | -104.1633
10.3392 | -104.1568
10.3429 | -104.1532
10.3466 | -104.1574
10.3504 | -104.1614
10.3541 | -104.1647
10.3578 | -104.1689
10.3615 | -104.1779
10.3653 | -104.1931
10.369 | -104.2106
10.3727 | -104.236
10.3764 | -104.2652
10.3802 | -104.2955
10.3839 | -104.327
10.3876 | -104.3627
10.3914 | -104.3999
10.3951 | -104.4397
10.3988 | -104.4801
10.4025 | -104.5195
10.4063 | -104.5559
10.41 | -104.5899
10.4137 | -104.6221
10.4174 | -104.6506
10.4212 | -104.6757
10.4249 | -104.6959
10.4286 | -104.7144
10.4324 | -104.7268
10.4361 | -104.7374
10.4398 | -104.744
10.4435 | -104.7451
10.4473 | -104.7384
10.451 | -104.7267
10.4547 | -104.709
10.4584 | -104.6888
10.4622 | -104.67
10.4659 | -104.6566
10.4696 | -104.6426
10.4734 | -104.6238
10.4771 | -104.608
10.4808 | -104.5928
10.4845 | -104.5763
10.4883 | -104.5639
10.492 | -104.5575
10.4957 | -104.5509
10.4994 | -104.5474
10.5032 | -104.5417
10.5069 | -104.5368
10.5106 | -104.5348
10.5144 | -104.5319
10.5181 | -104.5281
10.5218 | -104.5281
10.5255 | -104.5251
10.5293 | -104.5196
10.533 | -104.5141
10.5367 | -104.5097
10.5404 | -104.5039
10.5442 | -104.4976
10.5479 | -104.4881
10.5516 | -104.48
10.5553 | -104.4771
10.5591 | -104.4738
10.5628 | -104.4714
10.5665 | -104.4709
10.5703 | -104.471
10.574 | -104.4734
10.5777 | -104.4763
10.5814 | -104.483
10.5852 | -104.4936
10.5889 | -104.5079
10.5926 | -104.5275
10.5963 | -104.5504
10.6001 | -104.5752
10.6038 | -104.5981
10.6075 | -104.6246
10.6113 | -104.6512
10.615 | -104.6731
10.6187 | -104.694
10.6224 | -104.7154
10.6262 | -104.7347
10.6299 | -104.7522
10.6336 | -104.7689
10.6373 | -104.7855
10.6411 | -104.793
10.6448 | -104.791
10.6485 | -104.7863
10.6523 | -104.775
10.656 | -104.7675
10.6597 | -104.754
10.6634 | -104.7343
10.6672 | -104.713
10.6709 | -104.6918
10.6746 | -104.6657
10.6783 | -104.6412
10.6821 | -104.6231
10.6858 | -104.6024
10.6895 | -104.5807
10.6933 | -104.563
10.697 | -104.5481
10.7007 | -104.5379
10.7044 | -104.5316
10.7082 | -104.5259
10.7119 | -104.5252
10.7156 | -104.5303
10.7193 | -104.5373
10.7231 | -104.5466
10.7268 | -104.5615
10.7305 | -104.5801
10.7343 | -104.5975
10.738 | -104.613
10.7417 | -104.6264
10.7454 | -104.6377
10.7492 | -104.6453
10.7529 | -104.6396
10.7566 | -104.6417
10.7603 | -104.6425
10.7641 | -104.6392
10.7678 | -104.6285
10.7715 | -104.6172
10.7753 | -104.6045
10.779 | -104.5883
10.7827 | -104.5706
10.7864 | -104.5477
10.7902 | -104.5264
10.7939 | -104.5067
10.7976 | -104.4888
10.8013 | -104.4701
10.8051 | -104.4556
10.8088 | -104.446
10.8125 | -104.4364
10.8163 | -104.4313
10.82 | -104.4327
10.8237 | -104.435
10.8274 | -104.4387
10.8312 | -104.4487
10.8349 | -104.4603
10.8386 | -104.4722
10.8423 | -104.489
10.8461 | -104.511
10.8498 | -104.5316
10.8535 | -104.5524
10.8572 | -104.5744
10.861 | -104.5997
10.8647 | -104.6264
10.8684 | -104.652
10.8722 | -104.6742
10.8759 | -104.6951
10.8796 | -104.7124
10.8833 | -104.7255
10.8871 | -104.7362
10.8908 | -104.7403
10.8945 | -104.7446
10.8982 | -104.7448
10.902 | -104.7433
10.9057 | -104.7404
10.9094 | -104.7319
10.9132 | -104.7202
10.9169 | -104.7065
10.9206 | -104.6889
10.9243 | -104.6669
10.9281 | -104.648
10.9318 | -104.6269
10.9355 | -104.6053
10.9392 | -104.5882
10.943 | -104.5707
10.9467 | -104.5539
10.9504 | -104.5368
10.9542 | -104.5184
10.9579 | -104.5036
10.9616 | -104.493
10.9653 | -104.4814
10.9691 | -104.4705
10.9728 | -104.4623
10.9765 | -104.4575
10.9802 | -104.4504
10.984 | -104.4459
10.9877 | -104.4445
10.9914 | -104.4435
10.9952 | -104.4491
10.9989 | -104.4549
11.0026 | -104.4604
11.0063 | -104.4656
11.0101 | -104.4742
11.0138 | -104.4829
11.0175 | -104.49
11.0212 | -104.4982
11.025 | -104.5072
11.0287 | -104.5172
11.0324 | -104.5307
11.0362 | -104.5453
11.0399 | -104.5594
11.0436 | -104.5718
11.0473 | -104.5857
11.0511 | -104.6001
11.0548 | -104.6116
11.0585 | -104.6259
11.0622 | -104.6392
11.066 | -104.6526
11.0697 | -104.6638
11.0734 | -104.677
11.0772 | -104.6916
11.0809 | -104.7061
11.0846 | -104.7179
11.0883 | -104.7252
11.0921 | -104.7339
11.0958 | -104.742
11.0995 | -104.7461
11.1032 | -104.7481
11.107 | -104.7519
11.1107 | -104.7552
11.1144 | -104.76
11.1182 | -104.762
11.1219 | -104.7648
11.1256 | -104.7689
11.1293 | -104.7742
11.1331 | -104.777
11.1368 | -104.7791
11.1405 | -104.7837
11.1442 | -104.7896
11.148 | -104.7888
11.1517 | -104.7894
11.1554 | -104.7887
11.1592 | -104.7898
11.1629 | -104.7897
11.1666 | -104.7892
11.1703 | -104.7877
11.1741 | -104.7833
11.1778 | -104.7744
11.1815 | -104.7612
11.1852 | -104.7458
11.189 | -104.7283
11.1927 | -104.7066
11.1964 | -104.6822
11.2001 | -104.6582
11.2039 | -104.633
11.2076 | -104.6055
11.2113 | -104.5781
11.2151 | -104.5532
11.2188 | -104.5273
11.2225 | -104.5021
11.2262 | -104.4826
11.23 | -104.4703
11.2337 | -104.4609
11.2374 | -104.4545
11.2411 | -104.4499
11.2449 | -104.4512
11.2486 | -104.4573
11.2523 | -104.4588
11.2561 | -104.461
11.2598 | -104.4695
11.2635 | -104.4823
11.2672 | -104.4934
11.271 | -104.5039
11.2747 | -104.515
11.2784 | -104.5248
11.2821 | -104.534
11.2859 | -104.543
11.2896 | -104.5532
11.2933 | -104.564
11.2971 | -104.5759
11.3008 | -104.5899
11.3045 | -104.606
11.3082 | -104.6216
11.312 | -104.6426
11.3157 | -104.6613
11.3194 | -104.6785
11.3231 | -104.6967
11.3269 | -104.7178
11.3306 | -104.7377
11.3343 | -104.7565
11.3381 | -104.778
11.3418 | -104.8027
11.3455 | -104.82
11.3492 | -104.835
11.353 | -104.8471
11.3567 | -104.8577
11.3604 | -104.8679
11.3641 | -104.871
11.3679 | -104.8697
11.3716 | -104.8663
11.3753 | -104.8592
11.3791 | -104.847
11.3828 | -104.8334
11.3865 | -104.8162
11.3902 | -104.7939
11.394 | -104.7695
11.3977 | -104.739
11.4014 | -104.7113
11.4051 | -104.686
11.4089 | -104.6572
11.4126 | -104.6266
11.4163 | -104.6026
11.4201 | -104.5802
11.4238 | -104.5559
11.4275 | -104.5366
11.4312 | -104.5201
11.435 | -104.5071
11.4387 | -104.4971
11.4424 | -104.4887
11.4461 | -104.4864
11.4499 | -104.4878
11.4536 | -104.4911
11.4573 | -104.4944
11.4611 | -104.5024
11.4648 | -104.5122
11.4685 | -104.5206
11.4722 | -104.531
11.476 | -104.5438
11.4797 | -104.5543
11.4834 | -104.5644
11.4871 | -104.5746
11.4909 | -104.5838
11.4946 | -104.5882
11.4983 | -104.591
11.5021 | -104.5989
11.5058 | -104.6033
11.5095 | -104.6042
11.5132 | -104.6091
11.517 | -104.6138
11.5207 | -104.6198
11.5244 | -104.6249
11.5281 | -104.6282
11.5319 | -104.6335
11.5356 | -104.6407
11.5393 | -104.6481
11.543 | -104.6565
11.5468 | -104.665
11.5505 | -104.673
11.5542 | -104.6824
11.558 | -104.69
11.5617 | -104.6982
11.5654 | -104.7081
11.5691 | -104.7108
11.5729 | -104.7165
11.5766 | -104.7259
11.5803 | -104.7356
11.584 | -104.7425
11.5878 | -104.7504
11.5915 | -104.7591
11.5952 | -104.7674
11.599 | -104.7765
11.6027 | -104.7831
11.6064 | -104.7915
11.6101 | -104.8044
11.6139 | -104.8104
11.6176 | -104.8174
11.6213 | -104.8282
11.625 | -104.8395
11.6288 | -104.8478
11.6325 | -104.8614
11.6362 | -104.8794
11.64 | -104.8943
11.6437 | -104.911
11.6474 | -104.9275
11.6511 | -104.9459
11.6549 | -104.9669
11.6586 | -104.9862
11.6623 | -105.0083
11.666 | -105.0303
11.6698 | -105.0461
11.6735 | -105.0651
11.6772 | -105.0847
11.681 | -105.1023
11.6847 | -105.1127
11.6884 | -105.1212
11.6921 | -105.1298
11.6959 | -105.1321
11.6996 | -105.1352
11.7033 | -105.134
11.707 | -105.1273
11.7108 | -105.122
11.7145 | -105.1102
11.7182 | -105.0977
11.722 | -105.0802
11.7257 | -105.0637
11.7294 | -105.0442
11.7331 | -105.0236
11.7369 | -105.0031
11.7406 | -104.9773
11.7443 | -104.9495
11.748 | -104.9216
11.7518 | -104.894
11.7555 | -104.8642
11.7592 | -104.8385
11.763 | -104.8099
11.7667 | -104.7831
11.7704 | -104.7635
11.7741 | -104.7439
11.7779 | -104.7232
11.7816 | -104.7076
11.7853 | -104.6914
11.789 | -104.6776
11.7928 | -104.6657
11.7965 | -104.6564
11.8002 | -104.6481
11.804 | -104.6409
11.8077 | -104.6367
11.8114 | -104.6359
11.8151 | -104.6401
11.8189 | -104.6466
11.8226 | -104.6557
11.8263 | -104.6729
11.83 | -104.6847
11.8338 | -104.6979
11.8375 | -104.7128
11.8412 | -104.7277
11.845 | -104.7397
11.8487 | -104.7529
11.8524 | -104.7681
11.8561 | -104.7848
11.8599 | -104.7976
11.8636 | -104.8071
11.8673 | -104.8244
11.871 | -104.8389
11.8748 | -104.8507
11.8785 | -104.8655
11.8822 | -104.8764
11.8859 | -104.8882
11.8897 | -104.8979
11.8934 | -104.9093
11.8971 | -104.9169
11.9009 | -104.9221
11.9046 | -104.9265
11.9083 | -104.9293
11.912 | -104.9307
11.9158 | -104.9282
11.9195 | -104.9234
11.9232 | -104.9177
11.9269 | -104.9125
11.9307 | -104.9051
11.9344 | -104.8992
11.9381 | -104.8918
11.9419 | -104.8804
11.9456 | -104.8714
11.9493 | -104.8641
11.953 | -104.854
11.9568 | -104.8448
11.9605 | -104.8374
11.9642 | -104.8329
11.9679 | -104.8257
11.9717 | -104.8155
11.9754 | -104.8071
11.9791 | -104.801
11.9829 | -104.7932
11.9866 | -104.7885
11.9903 | -104.7831
11.994 | -104.774
11.9978 | -104.7682
12.0015 | -104.7684
12.0052 | -104.7699
12.0089 | -104.7735
12.0127 | -104.7766
12.0164 | -104.7805
12.0201 | -104.7865
12.0239 | -104.795
12.0276 | -104.8043
12.0313 | -104.814
12.035 | -104.8268
12.0388 | -104.8408
12.0425 | -104.8524
12.0462 | -104.8651
12.0499 | -104.8788
12.0537 | -104.8866
12.0574 | -104.8965
12.0611 | -104.91
12.0649 | -104.9184
12.0686 | -104.9261
12.0723 | -104.9381
12.076 | -104.9493
12.0798 | -104.9568
12.0835 | -104.9608
12.0872 | -104.9623
12.0909 | -104.9673
12.0947 | -104.9712
12.0984 | -104.9752
12.1021 | -104.9741
12.1059 | -104.9724
12.1096 | -104.9693
12.1133 | -104.9646
12.117 | -104.9594
12.1208 | -104.9509
12.1245 | -104.9453
12.1282 | -104.9404
12.1319 | -104.9352
12.1357 | -104.9269
12.1394 | -104.9209
12.1431 | -104.9192
12.1469 | -104.9201
12.1506 | -104.9213
12.1543 | -104.923
12.158 | -104.924
12.1618 | -104.9248
12.1655 | -104.929
12.1692 | -104.9353
12.1729 | -104.9392
12.1767 | -104.942
12.1804 | -104.9438
12.1841 | -104.9423
12.1878 | -104.9429
12.1916 | -104.9437
12.1953 | -104.9445
12.199 | -104.9416
12.2028 | -104.9391
12.2065 | -104.9404
12.2102 | -104.9417
12.2139 | -104.9434
12.2177 | -104.9448
12.2214 | -104.9468
12.2251 | -104.9538
12.2288 | -104.9625
12.2326 | -104.9697
12.2363 | -104.9764
12.24 | -104.9863
12.2438 | -104.993
12.2475 | -104.9963
12.2512 | -105.0024
12.2549 | -105.0137
12.2587 | -105.0248
12.2624 | -105.032
12.2661 | -105.0354
12.2698 | -105.0386
12.2736 | -105.0424
12.2773 | -105.0481
12.281 | -105.0516
12.2848 | -105.0569
12.2885 | -105.0631
12.2922 | -105.0702
12.2959 | -105.0774
12.2997 | -105.0797
12.3034 | -105.0841
12.3071 | -105.087
12.3108 | -105.0911
12.3146 | -105.0914
12.3183 | -105.0881
12.322 | -105.0883
12.3258 | -105.0891
12.3295 | -105.0891
12.3332 | -105.0915
12.3369 | -105.0922
12.3407 | -105.091
12.3444 | -105.0917
12.3481 | -105.0921
12.3518 | -105.0916
12.3556 | -105.093
12.3593 | -105.0967
12.363 | -105.0977
12.3668 | -105.1032
12.3705 | -105.1074
12.3742 | -105.1093
12.3779 | -105.1119
12.3817 | -105.1135
12.3854 | -105.1118
12.3891 | -105.1109
12.3928 | -105.1073
12.3966 | -105.1059
12.4003 | -105.1024
12.404 | -105.0972
12.4078 | -105.093
12.4115 | -105.0895
12.4152 | -105.083
12.4189 | -105.0745
12.4227 | -105.0681
12.4264 | -105.0579
12.4301 | -105.0459
12.4338 | -105.0331
12.4376 | -105.0233
12.4413 | -105.0136
12.445 | -105.0051
12.4488 | -104.9974
12.4525 | -104.9864
12.4562 | -104.9794
12.4599 | -104.9703
12.4637 | -104.9611
12.4674 | -104.9545
12.4711 | -104.9462
12.4748 | -104.94
12.4786 | -104.9361
12.4823 | -104.9305
12.486 | -104.9267
12.4898 | -104.9225
12.4935 | -104.9188
12.4972 | -104.9172
12.5009 | -104.9119
12.5047 | -104.9056
12.5084 | -104.9024
12.5121 | -104.8974
12.5158 | -104.8942
12.5196 | -104.8943
12.5233 | -104.8941
12.527 | -104.8945
12.5307 | -104.8965
12.5345 | -104.8998
12.5382 | -104.9051
12.5419 | -104.9133
12.5457 | -104.9182
12.5494 | -104.9268
12.5531 | -104.9396
12.5568 | -104.9541
12.5606 | -104.9703
12.5643 | -104.9899
12.568 | -105.0142
12.5717 | -105.0391
12.5755 | -105.0656
12.5792 | -105.0924
12.5829 | -105.1227
12.5867 | -105.1521
12.5904 | -105.1821
12.5941 | -105.2109
12.5978 | -105.24
12.6016 | -105.2693
12.6053 | -105.2957
12.609 | -105.3198
12.6127 | -105.3404
12.6165 | -105.3577
12.6202 | -105.3719
12.6239 | -105.3865
12.6277 | -105.3981
12.6314 | -105.4057
12.6351 | -105.4091
12.6388 | -105.4101
12.6426 | -105.4128
12.6463 | -105.4124
12.65 | -105.4071
12.6537 | -105.4004
12.6575 | -105.3915
12.6612 | -105.3803
12.6649 | -105.3699
12.6687 | -105.3558
12.6724 | -105.3407
12.6761 | -105.3237
12.6798 | -105.3048
12.6836 | -105.2857
12.6873 | -105.2681
12.691 | -105.2472
12.6947 | -105.2291
12.6985 | -105.2143
12.7022 | -105.1993
12.7059 | -105.185
12.7097 | -105.1735
12.7134 | -105.1657
12.7171 | -105.1582
12.7208 | -105.1517
12.7246 | -105.1487
12.7283 | -105.1515
12.732 | -105.1562
12.7357 | -105.1598
12.7395 | -105.1664
12.7432 | -105.174
12.7469 | -105.1836
12.7507 | -105.1927
12.7544 | -105.2062
12.7581 | -105.2208
12.7618 | -105.2327
12.7656 | -105.2444
12.7693 | -105.2578
12.773 | -105.2716
12.7767 | -105.2866
12.7805 | -105.2997
12.7842 | -105.3111
12.7879 | -105.3208
12.7917 | -105.3279
12.7954 | -105.3333
12.7991 | -105.3369
12.8028 | -105.338
12.8066 | -105.3376
12.8103 | -105.333
12.814 | -105.3289
12.8177 | -105.3262
12.8215 | -105.3213
12.8252 | -105.316
12.8289 | -105.3097
12.8326 | -105.3054
12.8364 | -105.3021
12.8401 | -105.2984
12.8438 | -105.2939
12.8476 | -105.288
12.8513 | -105.2846
12.855 | -105.282
12.8587 | -105.2807
12.8625 | -105.2796
12.8662 | -105.2789
12.8699 | -105.2752
12.8736 | -105.2741
12.8774 | -105.2752
12.8811 | -105.2751
12.8848 | -105.2744
12.8886 | -105.2729
12.8923 | -105.2738
12.896 | -105.2739
12.8997 | -105.274
12.9035 | -105.2707
12.9072 | -105.2711
12.9109 | -105.2686
12.9146 | -105.2661
12.9184 | -105.2646
12.9221 | -105.2652
12.9258 | -105.2649
12.9296 | -105.2641
12.9333 | -105.2627
12.937 | -105.2621
12.9407 | -105.2613
12.9445 | -105.2663
12.9482 | -105.273
12.9519 | -105.2786
12.9556 | -105.2831
12.9594 | -105.2896
12.9631 | -105.2944
12.9668 | -105.3054
12.9706 | -105.319
12.9743 | -105.3317
12.978 | -105.3489
12.9817 | -105.3643
12.9855 | -105.3778
12.9892 | -105.3919
12.9929 | -105.4059
12.9966 | -105.4194
13.0004 | -105.431
13.0041 | -105.44
13.0078 | -105.4449
13.0116 | -105.4507
13.0153 | -105.4522
13.019 | -105.4489
13.0227 | -105.4451
13.0265 | -105.4381
13.0302 | -105.4313
13.0339 | -105.4225
13.0376 | -105.4093
13.0414 | -105.3927
13.0451 | -105.3763
13.0488 | -105.3602
13.0526 | -105.3425
13.0563 | -105.3236
13.06 | -105.3076
13.0637 | -105.2893
13.0675 | -105.2702
13.0712 | -105.2526
13.0749 | -105.2408
13.0786 | -105.2273
13.0824 | -105.2127
13.0861 | -105.2021
13.0898 | -105.195
13.0936 | -105.1883
13.0973 | -105.1842
13.101 | -105.1826
13.1047 | -105.1817
13.1085 | -105.1832
13.1122 | -105.1847
13.1159 | -105.1887
13.1196 | -105.1933
13.1234 | -105.1984
13.1271 | -105.2018
13.1308 | -105.2003
13.1346 | -105.1985
13.1383 | -105.1998
13.142 | -105.2054
13.1457 | -105.2112
13.1495 | -105.2122
13.1532 | -105.2121
13.1569 | -105.2099
13.1606 | -105.2082
13.1644 | -105.2045
13.1681 | -105.202
13.1718 | -105.1999
13.1755 | -105.1971
13.1793 | -105.1932
13.183 | -105.1884
13.1867 | -105.1888
13.1905 | -105.1871
13.1942 | -105.184
13.1979 | -105.1826
13.2016 | -105.1799
13.2054 | -105.1812
13.2091 | -105.1851
13.2128 | -105.1863
13.2165 | -105.1873
13.2203 | -105.1857
13.224 | -105.185
13.2277 | -105.1859
13.2315 | -105.1864
13.2352 | -105.1858
13.2389 | -105.1822
13.2426 | -105.1779
13.2464 | -105.1744
13.2501 | -105.1638
13.2538 | -105.1528
13.2575 | -105.1439
13.2613 | -105.1342
13.265 | -105.1236
13.2687 | -105.1142
13.2725 | -105.1028
13.2762 | -105.0914
13.2799 | -105.0791
13.2836 | -105.0692
13.2874 | -105.0623
13.2911 | -105.0555
13.2948 | -105.0493
13.2985 | -105.045
13.3023 | -105.0453
13.306 | -105.0465
13.3097 | -105.0489
13.3135 | -105.0531
13.3172 | -105.0599
13.3209 | -105.0716
13.3246 | -105.0813
13.3284 | -105.0903
13.3321 | -105.0982
13.3358 | -105.106
13.3395 | -105.1175
13.3433 | -105.133
13.347 | -105.1475
13.3507 | -105.162
13.3545 | -105.1783
13.3582 | -105.1916
13.3619 | -105.2047
13.3656 | -105.2177
13.3694 | -105.228
13.3731 | -105.2391
13.3768 | -105.2473
13.3805 | -105.2582
13.3843 | -105.2716
13.388 | -105.2827
13.3917 | -105.2914
13.3955 | -105.2969
13.3992 | -105.3021
13.4029 | -105.3084
13.4066 | -105.3154
13.4104 | -105.3235
13.4141 | -105.3312
13.4178 | -105.3333
13.4215 | -105.337
13.4253 | -105.3451
13.429 | -105.3546
13.4327 | -105.3634
13.4365 | -105.374
13.4402 | -105.3813
13.4439 | -105.3913
13.4476 | -105.4012
13.4514 | -105.4068
13.4551 | -105.4116
13.4588 | -105.4152
13.4625 | -105.4154
13.4663 | -105.415
13.47 | -105.4124
13.4737 | -105.4106
13.4775 | -105.404
13.4812 | -105.3911
13.4849 | -105.3794
13.4886 | -105.3656
13.4924 | -105.3464
13.4961 | -105.3259
13.4998 | -105.3092
13.5035 | -105.2886
13.5073 | -105.2678
13.511 | -105.2493
13.5147 | -105.2301
13.5184 | -105.2122
13.5222 | -105.1949
13.5259 | -105.1789
13.5296 | -105.1662
13.5334 | -105.1516
13.5371 | -105.1371
13.5408 | -105.1216
13.5445 | -105.109
13.5483 | -105.102
13.552 | -105.0971
13.5557 | -105.0926
13.5594 | -105.0921
13.5632 | -105.0923
13.5669 | -105.0916
13.5706 | -105.0912
13.5744 | -105.0915
13.5781 | -105.0951
13.5818 | -105.1012
13.5855 | -105.1106
13.5893 | -105.1227
13.593 | -105.1337
13.5967 | -105.1448
13.6004 | -105.1574
13.6042 | -105.1698
13.6079 | -105.1807
13.6116 | -105.195
13.6154 | -105.2099
13.6191 | -105.2196
13.6228 | -105.2294
13.6265 | -105.2415
13.6303 | -105.2508
13.634 | -105.2601
13.6377 | -105.2707
13.6414 | -105.2788
13.6452 | -105.2857
13.6489 | -105.2903
13.6526 | -105.2947
13.6564 | -105.3009
13.6601 | -105.3102
13.6638 | -105.3194
13.6675 | -105.3259
13.6713 | -105.3283
13.675 | -105.3323
13.6787 | -105.3376
13.6824 | -105.3422
13.6862 | -105.3458
13.6899 | -105.3503
13.6936 | -105.3523
13.6974 | -105.3516
13.7011 | -105.3531
13.7048 | -105.3538
13.7085 | -105.3525
13.7123 | -105.3499
13.716 | -105.3467
13.7197 | -105.3424
13.7234 | -105.3392
13.7272 | -105.3354
13.7309 | -105.3352
13.7346 | -105.3325
13.7384 | -105.3336
13.7421 | -105.3348
13.7458 | -105.3339
13.7495 | -105.3325
13.7533 | -105.3366
13.757 | -105.3376
13.7607 | -105.3405
13.7644 | -105.3453
13.7682 | -105.3485
13.7719 | -105.3532
13.7756 | -105.3531
13.7794 | -105.3503
13.7831 | -105.3484
13.7868 | -105.3479
13.7905 | -105.3463
13.7943 | -105.3431
13.798 | -105.3383
13.8017 | -105.3314
13.8054 | -105.3236
13.8092 | -105.3156
13.8129 | -105.3028
13.8166 | -105.2904
13.8203 | -105.2782
13.8241 | -105.2633
13.8278 | -105.2472
13.8315 | -105.2315
13.8353 | -105.2134
13.839 | -105.1938
13.8427 | -105.1772
13.8464 | -105.1633
13.8502 | -105.149
13.8539 | -105.1297
13.8576 | -105.1128
13.8613 | -105.0982
13.8651 | -105.0872
13.8688 | -105.0774
13.8725 | -105.069
13.8763 | -105.0622
13.88 | -105.056
13.8837 | -105.0494
13.8874 | -105.0473
13.8912 | -105.0471
13.8949 | -105.0443
13.8986 | -105.0439
13.9023 | -105.0455
13.9061 | -105.0502
13.9098 | -105.0581
13.9135 | -105.0666
13.9173 | -105.0785
13.921 | -105.0909
13.9247 | -105.104
13.9284 | -105.1177
13.9322 | -105.132
13.9359 | -105.1472
13.9396 | -105.1615
13.9433 | -105.1765
13.9471 | -105.1943
13.9508 | -105.2094
13.9545 | -105.2278
13.9583 | -105.2475
13.962 | -105.2597
13.9657 | -105.2677
13.9694 | -105.2752
13.9732 | -105.2852
13.9769 | -105.2874
13.9806 | -105.2995
13.9843 | -105.3064
13.9881 | -105.3113
13.9918 | -105.3166
13.9955 | -105.3202
13.9993 | -105.3211
14.003 | -105.3199
14.0067 | -105.3158
14.0104 | -105.3103
14.0142 | -105.3042
14.0179 | -105.2963
14.0216 | -105.2838
14.0253 | -105.2742
14.0291 | -105.2652
14.0328 | -105.2568
14.0365 | -105.2478
14.0403 | -105.2412
14.044 | -105.2332
14.0477 | -105.2265
14.0514 | -105.221
14.0552 | -105.2184
14.0589 | -105.2163
14.0626 | -105.2151
14.0663 | -105.2138
14.0701 | -105.2182
14.0738 | -105.226
14.0775 | -105.2338
14.0813 | -105.2419
14.085 | -105.2536
14.0887 | -105.2644
14.0924 | -105.278
14.0962 | -105.2917
14.0999 | -105.302
14.1036 | -105.3107
14.1073 | -105.3161
14.1111 | -105.3194
14.1148 | -105.3234
14.1185 | -105.3251
14.1223 | -105.3266
14.126 | -105.3261
14.1297 | -105.3208
14.1334 | -105.3144
14.1372 | -105.3043
14.1409 | -105.2942
14.1446 | -105.285
14.1483 | -105.2747
14.1521 | -105.2637
14.1558 | -105.2507
14.1595 | -105.2357
14.1632 | -105.2263
14.167 | -105.2202
14.1707 | -105.2107
14.1744 | -105.1978
14.1782 | -105.1861
14.1819 | -105.1753
14.1856 | -105.165
14.1893 | -105.1581
14.1931 | -105.1552
14.1968 | -105.155
14.2005 | -105.1568
14.2042 | -105.1578
14.208 | -105.1632
14.2117 | -105.1709
14.2154 | -105.1786
14.2192 | -105.1869
14.2229 | -105.1963
14.2266 | -105.2103
14.2303 | -105.2232
14.2341 | -105.2358
14.2378 | -105.2497
14.2415 | -105.2647
14.2452 | -105.2811
14.249 | -105.296
14.2527 | -105.3107
14.2564 | -105.3261
14.2602 | -105.3414
14.2639 | -105.3554
14.2676 | -105.3699
14.2713 | -105.3852
14.2751 | -105.395
14.2788 | -105.4071
14.2825 | -105.42
14.2862 | -105.4345
14.29 | -105.4477
14.2937 | -105.4599
14.2974 | -105.4695
14.3012 | -105.4795
14.3049 | -105.4937
14.3086 | -105.5046
14.3123 | -105.5142
14.3161 | -105.5249
14.3198 | -105.5325
14.3235 | -105.5375
14.3272 | -105.5391
14.331 | -105.5435
14.3347 | -105.5484
14.3384 | -105.5535
14.3422 | -105.5589
14.3459 | -105.5585
14.3496 | -105.5557
14.3533 | -105.5524
14.3571 | -105.5482
14.3608 | -105.543
14.3645 | -105.5376
14.3682 | -105.5292
14.372 | -105.5195
14.3757 | -105.5107
14.3794 | -105.502
14.3832 | -105.4911
14.3869 | -105.4781
14.3906 | -105.4652
14.3943 | -105.4546
14.3981 | -105.4398
14.4018 | -105.4239
14.4055 | -105.4128
14.4092 | -105.4062
14.413 | -105.3978
14.4167 | -105.3894
14.4204 | -105.3812
14.4242 | -105.3758
14.4279 | -105.3713
14.4316 | -105.3652
14.4353 | -105.3613
14.4391 | -105.3575
14.4428 | -105.3522
14.4465 | -105.3486
14.4502 | -105.3443
14.454 | -105.342
14.4577 | -105.3407
14.4614 | -105.337
14.4652 | -105.3334
14.4689 | -105.3305
14.4726 | -105.3281
14.4763 | -105.3246
14.4801 | -105.3227
14.4838 | -105.3163
14.4875 | -105.3084
14.4912 | -105.2992
14.495 | -105.2911
14.4987 | -105.2826
14.5024 | -105.2742
14.5061 | -105.269
14.5099 | -105.2619
14.5136 | -105.2562
14.5173 | -105.2482
14.5211 | -105.2424
14.5248 | -105.2335
14.5285 | -105.2266
14.5322 | -105.2204
14.536 | -105.2131
14.5397 | -105.2052
14.5434 | -105.1967
14.5471 | -105.1877
14.5509 | -105.1798
14.5546 | -105.175
14.5583 | -105.1719
14.5621 | -105.1683
14.5658 | -105.1679
14.5695 | -105.1685
14.5732 | -105.1706
14.577 | -105.1719
14.5807 | -105.1748
14.5844 | -105.1798
14.5881 | -105.1864
14.5919 | -105.1952
14.5956 | -105.2051
14.5993 | -105.2156
14.6031 | -105.2263
14.6068 | -105.2397
14.6105 | -105.2555
14.6142 | -105.2686
14.618 | -105.2808
14.6217 | -105.2953
14.6254 | -105.3105
14.6291 | -105.3241
14.6329 | -105.3348
14.6366 | -105.3463
14.6403 | -105.3596
14.6441 | -105.3694
14.6478 | -105.3787
14.6515 | -105.3891
14.6552 | -105.3996
14.659 | -105.4086
14.6627 | -105.4157
14.6664 | -105.4226
14.6701 | -105.4319
14.6739 | -105.4392
14.6776 | -105.4428
14.6813 | -105.4457
14.6851 | -105.4493
14.6888 | -105.4526
14.6925 | -105.4553
14.6962 | -105.4579
14.7 | -105.46
14.7037 | -105.4618
14.7074 | -105.4622
14.7111 | -105.4616
14.7149 | -105.4585
14.7186 | -105.4564
14.7223 | -105.4544
14.7261 | -105.4564
14.7298 | -105.4582
14.7335 | -105.4624
14.7372 | -105.4651
14.741 | -105.4661
14.7447 | -105.4683
14.7484 | -105.4734
14.7521 | -105.4792
14.7559 | -105.4841
14.7596 | -105.4916
14.7633 | -105.4993
14.7671 | -105.5075
14.7708 | -105.5142
14.7745 | -105.5195
14.7782 | -105.5214
14.782 | -105.5232
14.7857 | -105.5272
14.7894 | -105.532
14.7931 | -105.5343
14.7969 | -105.5311
14.8006 | -105.5284
14.8043 | -105.5237
14.8081 | -105.5199
14.8118 | -105.5149
14.8155 | -105.5057
14.8192 | -105.4961
14.823 | -105.4858
14.8267 | -105.4779
14.8304 | -105.4717
14.8341 | -105.4634
14.8379 | -105.4556
14.8416 | -105.4487
14.8453 | -105.4427
14.8491 | -105.435
14.8528 | -105.429
14.8565 | -105.4245
14.8602 | -105.4219
14.864 | -105.419
14.8677 | -105.4149
14.8714 | -105.4126
14.8751 | -105.4129
14.8789 | -105.4132
14.8826 | -105.411
14.8863 | -105.4089
14.89 | -105.4077
14.8938 | -105.4056
14.8975 | -105.4037
14.9012 | -105.4029
14.905 | -105.4025
14.9087 | -105.4042
14.9124 | -105.4067
14.9161 | -105.4073
14.9199 | -105.4089
14.9236 | -105.4108
14.9273 | -105.4122
14.931 | -105.4125
14.9348 | -105.4155
14.9385 | -105.416
14.9422 | -105.4142
14.946 | -105.4132
14.9497 | -105.4157
14.9534 | -105.419
14.9571 | -105.4211
14.9609 | -105.4218
14.9646 | -105.4189
14.9683 | -105.4155
14.972 | -105.407
14.9758 | -105.3976
14.9795 | -105.3909
14.9832 | -105.3821
14.987 | -105.3737
14.9907 | -105.3678
14.9944 | -105.3607
14.9981 | -105.3517
15.0019 | -105.3413
15.0056 | -105.3308
15.0093 | -105.322
15.013 | -105.3109
15.0168 | -105.2998
15.0205 | -105.2868
15.0242 | -105.2758
15.028 | -105.2662
15.0317 | -105.2555
15.0354 | -105.2436
15.0391 | -105.2281
15.0429 | -105.2117
15.0466 | -105.1951
15.0503 | -105.1794
15.054 | -105.1692
15.0578 | -105.1582
15.0615 | -105.1504
15.0652 | -105.1444
15.069 | -105.1347
15.0727 | -105.1262
15.0764 | -105.1174
15.0801 | -105.1084
15.0839 | -105.1024
15.0876 | -105.0991
15.0913 | -105.0926
15.095 | -105.089
15.0988 | -105.0899
15.1025 | -105.0923
15.1062 | -105.0937
15.11 | -105.0952
15.1137 | -105.0963
15.1174 | -105.0951
15.1211 | -105.0959
15.1249 | -105.0983
15.1286 | -105.101
15.1323 | -105.1056
15.136 | -105.1118
15.1398 | -105.1168
15.1435 | -105.125
15.1472 | -105.1337
15.1509 | -105.1421
15.1547 | -105.1527
15.1584 | -105.1647
15.1621 | -105.1758
15.1659 | -105.187
15.1696 | -105.1994
15.1733 | -105.2139
15.177 | -105.2277
15.1808 | -105.2408
15.1845 | -105.2566
15.1882 | -105.2738
15.1919 | -105.2888
15.1957 | -105.3022
15.1994 | -105.3181
15.2031 | -105.334
15.2069 | -105.3477
15.2106 | -105.3591
15.2143 | -105.369
15.218 | -105.3804
15.2218 | -105.3907
15.2255 | -105.3966
15.2292 | -105.4042
15.2329 | -105.4123
15.2367 | -105.4201
15.2404 | -105.4233
15.2441 | -105.4265
15.2479 | -105.4257
15.2516 | -105.4248
15.2553 | -105.4263
15.259 | -105.4266
15.2628 | -105.4277
15.2665 | -105.4271
15.2702 | -105.4208
15.2739 | -105.4152
15.2777 | -105.4087
15.2814 | -105.3998
15.2851 | -105.3895
15.2889 | -105.382
15.2926 | -105.3745
15.2963 | -105.3689
15.3 | -105.3635
15.3038 | -105.3533
15.3075 | -105.3432
15.3112 | -105.3381
15.3149 | -105.3297
15.3187 | -105.3198
15.3224 | -105.3126
15.3261 | -105.3072
15.3299 | -105.3024
15.3336 | -105.3002
15.3373 | -105.2984
15.341 | -105.2936
15.3448 | -105.2909
15.3485 | -105.2886
15.3522 | -105.2872
15.3559 | -105.2856
15.3597 | -105.2879
15.3634 | -105.2877
15.3671 | -105.2832
15.3709 | -105.2806
15.3746 | -105.2805
15.3783 | -105.2799
15.382 | -105.2813
15.3858 | -105.2812
15.3895 | -105.2814
15.3932 | -105.2814
15.3969 | -105.2805
15.4007 | -105.2795
15.4044 | -105.2785
15.4081 | -105.278
15.4119 | -105.2765
15.4156 | -105.2708
15.4193 | -105.2656
15.423 | -105.2608
15.4268 | -105.2532
15.4305 | -105.2478
15.4342 | -105.2452
15.4379 | -105.2418
15.4417 | -105.2393
15.4454 | -105.2416
15.4491 | -105.2375
15.4529 | -105.2302
15.4566 | -105.2268
15.4603 | -105.224
15.464 | -105.2217
15.4678 | -105.2227
15.4715 | -105.2211
15.4752 | -105.2172
15.4789 | -105.2133
15.4827 | -105.2093
15.4864 | -105.2029
15.4901 | -105.2023
15.4939 | -105.203
15.4976 | -105.2039
15.5013 | -105.2028
15.505 | -105.2029
15.5088 | -105.2028
15.5125 | -105.2035
15.5162 | -105.2032
15.5199 | -105.2035
15.5237 | -105.2066
15.5274 | -105.2076
15.5311 | -105.2065
15.5348 | -105.2065
15.5386 | -105.2082
15.5423 | -105.2105
15.546 | -105.2145
15.5498 | -105.2165
15.5535 | -105.2167
15.5572 | -105.2203
15.5609 | -105.2239
15.5647 | -105.2266
15.5684 | -105.2299
15.5721 | -105.2352
15.5758 | -105.2415
15.5796 | -105.2488
15.5833 | -105.2581
15.587 | -105.2727
15.5908 | -105.2933
15.5945 | -105.306
15.5982 | -105.3181
15.6019 | -105.3321
15.6057 | -105.3466
15.6094 | -105.3628
15.6131 | -105.3776
15.6168 | -105.3931
15.6206 | -105.4139
15.6243 | -105.4383
15.628 | -105.4603
15.6318 | -105.4796
15.6355 | -105.499
15.6392 | -105.5226
15.6429 | -105.5404
15.6467 | -105.556
15.6504 | -105.5753
15.6541 | -105.5906
15.6578 | -105.6008
15.6616 | -105.6091
15.6653 | -105.6156
15.669 | -105.619
15.6728 | -105.6232
15.6765 | -105.6259
15.6802 | -105.6231
15.6839 | -105.6188
15.6877 | -105.6114
15.6914 | -105.5988
15.6951 | -105.5868
15.6988 | -105.5734
15.7026 | -105.5572
15.7063 | -105.5401
15.71 | -105.5229
15.7138 | -105.5017
15.7175 | -105.4791
15.7212 | -105.4588
15.7249 | -105.4399
15.7287 | -105.4201
15.7324 | -105.4009
15.7361 | -105.3798
15.7398 | -105.3618
15.7436 | -105.3448
15.7473 | -105.3307
15.751 | -105.319
15.7548 | -105.3097
15.7585 | -105.3027
15.7622 | -105.2953
15.7659 | -105.2879
15.7697 | -105.2827
15.7734 | -105.2835
15.7771 | -105.2866
15.7808 | -105.2939
15.7846 | -105.3023
15.7883 | -105.3114
15.792 | -105.3253
15.7958 | -105.3372
15.7995 | -105.3497
15.8032 | -105.3635
15.8069 | -105.3794
15.8107 | -105.3952
15.8144 | -105.4109
15.8181 | -105.4299
15.8218 | -105.4477
15.8256 | -105.4635
15.8293 | -105.481
15.833 | -105.496
15.8368 | -105.5098
15.8405 | -105.5222
15.8442 | -105.5322
15.8479 | -105.542
15.8517 | -105.5499
15.8554 | -105.5584
15.8591 | -105.5683
15.8628 | -105.5762
15.8666 | -105.5804
15.8703 | -105.5871
15.874 | -105.5926
15.8777 | -105.5948
15.8815 | -105.5973
15.8852 | -105.601
15.8889 | -105.6014
15.8927 | -105.603
15.8964 | -105.6061
15.9001 | -105.6094
15.9038 | -105.6119
15.9076 | -105.6133
15.9113 | -105.6144
15.915 | -105.6148
15.9187 | -105.6184
15.9225 | -105.6188
15.9262 | -105.618
15.9299 | -105.6184
15.9337 | -105.6186
15.9374 | -105.6188
15.9411 | -105.6209
15.9448 | -105.6208
15.9486 | -105.6214
15.9523 | -105.623
15.956 | -105.6242
15.9597 | -105.6253
15.9635 | -105.6266
15.9672 | -105.6242
15.9709 | -105.6195
15.9747 | -105.6144
15.9784 | -105.6111
15.9821 | -105.6062
15.9858 | -105.6002
15.9896 | -105.595
15.9933 | -105.5906
15.997 | -105.5842
16.0007 | -105.5741
16.0045 | -105.5645
16.0082 | -105.5514
16.0119 | -105.5417
16.0157 | -105.5306
16.0194 | -105.5186
16.0231 | -105.5088
16.0268 | -105.5008
16.0306 | -105.488
16.0343 | -105.4756
16.038 | -105.4646
16.0417 | -105.4532
16.0455 | -105.4424
16.0492 | -105.435
16.0529 | -105.4258
16.0567 | -105.415
16.0604 | -105.4046
16.0641 | -105.3947
16.0678 | -105.3868
16.0716 | -105.3797
16.0753 | -105.3708
16.079 | -105.3592
16.0827 | -105.3455
16.0865 | -105.3347
16.0902 | -105.3271
16.0939 | -105.3211
16.0977 | -105.3134
16.1014 | -105.3077
16.1051 | -105.303
16.1088 | -105.2999
16.1126 | -105.2942
16.1163 | -105.2926
16.12 | -105.29
16.1237 | -105.288
16.1275 | -105.2864
16.1312 | -105.2877
16.1349 | -105.2872
16.1387 | -105.2898
16.1424 | -105.2935
16.1461 | -105.2963
16.1498 | -105.2992
16.1536 | -105.3046
16.1573 | -105.3105
16.161 | -105.3189
16.1647 | -105.3257
16.1685 | -105.3334
16.1722 | -105.3407
16.1759 | -105.35
16.1796 | -105.3613
16.1834 | -105.3733
16.1871 | -105.385
16.1908 | -105.3978
16.1946 | -105.412
16.1983 | -105.4292
16.202 | -105.4475
16.2057 | -105.4637
16.2095 | -105.4765
16.2132 | -105.493
16.2169 | -105.5108
16.2206 | -105.5275
16.2244 | -105.5405
16.2281 | -105.5543
16.2318 | -105.5701
16.2356 | -105.5841
16.2393 | -105.6006
16.243 | -105.6153
16.2467 | -105.6282
16.2505 | -105.6389
16.2542 | -105.6513
16.2579 | -105.6615
16.2616 | -105.6733
16.2654 | -105.6842
16.2691 | -105.693
16.2728 | -105.7012
16.2766 | -105.7062
16.2803 | -105.7112
16.284 | -105.7178
16.2877 | -105.7252
16.2915 | -105.7319
16.2952 | -105.7387
16.2989 | -105.7431
16.3026 | -105.7467
16.3064 | -105.7478
16.3101 | -105.7516
16.3138 | -105.755
16.3176 | -105.7546
16.3213 | -105.7507
16.325 | -105.7452
16.3287 | -105.7388
16.3325 | -105.7332
16.3362 | -105.7264
16.3399 | -105.7161
16.3436 | -105.7069
16.3474 | -105.6963
16.3511 | -105.685
16.3548 | -105.6738
16.3586 | -105.6603
16.3623 | -105.6484
16.366 | -105.6362
16.3697 | -105.6222
16.3735 | -105.605
16.3772 | -105.586
16.3809 | -105.57
16.3846 | -105.5538
16.3884 | -105.5358
16.3921 | -105.5202
16.3958 | -105.5016
16.3996 | -105.4837
16.4033 | -105.4683
16.407 | -105.4531
16.4107 | -105.4358
16.4145 | -105.4173
16.4182 | -105.4032
16.4219 | -105.3926
16.4256 | -105.3809
16.4294 | -105.3703
16.4331 | -105.3595
16.4368 | -105.3514
16.4406 | -105.3444
16.4443 | -105.3408
16.448 | -105.3373
16.4517 | -105.3344
16.4555 | -105.3335
16.4592 | -105.3378
16.4629 | -105.3438
16.4666 | -105.3483
16.4704 | -105.3516
16.4741 | -105.3553
16.4778 | -105.3603
16.4816 | -105.3701
16.4853 | -105.3812
16.489 | -105.392
16.4927 | -105.4021
16.4965 | -105.412
16.5002 | -105.4211
16.5039 | -105.4324
16.5076 | -105.4437
16.5114 | -105.4547
16.5151 | -105.463
16.5188 | -105.4715
16.5225 | -105.4791
16.5263 | -105.4855
16.53 | -105.492
16.5337 | -105.4973
16.5375 | -105.5016
16.5412 | -105.5072
16.5449 | -105.5119
16.5486 | -105.5164
16.5524 | -105.5195
16.5561 | -105.5232
16.5598 | -105.5256
16.5635 | -105.5281
16.5673 | -105.5271
16.571 | -105.527
16.5747 | -105.5295
16.5785 | -105.5321
16.5822 | -105.5324
16.5859 | -105.5328
16.5896 | -105.5335
16.5934 | -105.5385
16.5971 | -105.5426
16.6008 | -105.5453
16.6045 | -105.5495
16.6083 | -105.554
16.612 | -105.5586
16.6157 | -105.5616
16.6195 | -105.566
16.6232 | -105.5725
16.6269 | -105.5764
16.6306 | -105.5819
16.6344 | -105.5887
16.6381 | -105.5937
16.6418 | -105.5943
16.6455 | -105.5987
16.6493 | -105.6059
16.653 | -105.6099
16.6567 | -105.6145
16.6605 | -105.6206
16.6642 | -105.6275
16.6679 | -105.6281
16.6716 | -105.6312
16.6754 | -105.6358
16.6791 | -105.6428
16.6828 | -105.6464
16.6865 | -105.6497
16.6903 | -105.6557
16.694 | -105.6615
16.6977 | -105.6646
16.7015 | -105.6686
16.7052 | -105.6755
16.7089 | -105.6833
16.7126 | -105.6893
16.7164 | -105.6961
16.7201 | -105.6999
16.7238 | -105.7043
16.7275 | -105.7078
16.7313 | -105.7127
16.735 | -105.7188
16.7387 | -105.7247
16.7425 | -105.7287
16.7462 | -105.7323
16.7499 | -105.7381
16.7536 | -105.7451
16.7574 | -105.751
16.7611 | -105.7549
16.7648 | -105.7603
16.7685 | -105.7659
16.7723 | -105.7711
16.776 | -105.7769
16.7797 | -105.7834
16.7835 | -105.7879
16.7872 | -105.7944
16.7909 | -105.8005
16.7946 | -105.8056
16.7984 | -105.8117
16.8021 | -105.8197
16.8058 | -105.8253
16.8095 | -105.8316
16.8133 | -105.8375
16.817 | -105.843
16.8207 | -105.8507
16.8244 | -105.8562
16.8282 | -105.859
16.8319 | -105.8606
16.8356 | -105.8621
16.8394 | -105.865
16.8431 | -105.8674
16.8468 | -105.8696
16.8505 | -105.8727
16.8543 | -105.874
16.858 | -105.8729
16.8617 | -105.8684
16.8654 | -105.8641
16.8692 | -105.8599
16.8729 | -105.8573
16.8766 | -105.8548
16.8804 | -105.8558
16.8841 | -105.8538
16.8878 | -105.849
16.8915 | -105.8463
16.8953 | -105.8435
16.899 | -105.8417
16.9027 | -105.8407
16.9064 | -105.8422
16.9102 | -105.8478
16.9139 | -105.8499
16.9176 | -105.8515
16.9214 | -105.852
16.9251 | -105.8546
16.9288 | -105.8585
16.9325 | -105.8608
16.9363 | -105.8638
16.94 | -105.8696
16.9437 | -105.8769
16.9474 | -105.8825
16.9512 | -105.8906
16.9549 | -105.8994
16.9586 | -105.9034
16.9624 | -105.9081
16.9661 | -105.9124
16.9698 | -105.9179
16.9735 | -105.9251
16.9773 | -105.9306
16.981 | -105.9359
16.9847 | -105.9409
16.9884 | -105.9466
16.9922 | -105.951
16.9959 | -105.9549
16.9996 | -105.9568
17.0034 | -105.9594
17.0071 | -105.9619
17.0108 | -105.9636
17.0145 | -105.9664
17.0183 | -105.9672
17.022 | -105.9662
17.0257 | -105.9657
17.0294 | -105.964
17.0332 | -105.9613
17.0369 | -105.955
17.0406 | -105.9465
17.0444 | -105.9376
17.0481 | -105.9299
17.0518 | -105.9217
17.0555 | -105.9122
17.0593 | -105.9009
17.063 | -105.8908
17.0667 | -105.8828
17.0704 | -105.8752
17.0742 | -105.8661
17.0779 | -105.8569
17.0816 | -105.8474
17.0854 | -105.8369
17.0891 | -105.8266
17.0928 | -105.8168
17.0965 | -105.8075
17.1003 | -105.798
17.104 | -105.7904
17.1077 | -105.7833
17.1114 | -105.7773
17.1152 | -105.7699
17.1189 | -105.7619
17.1226 | -105.7543
17.1264 | -105.7508
17.1301 | -105.7472
17.1338 | -105.7438
17.1375 | -105.7417
17.1413 | -105.7395
17.145 | -105.7381
17.1487 | -105.7387
17.1524 | -105.7407
17.1562 | -105.746
17.1599 | -105.75
17.1636 | -105.7561
17.1673 | -105.7643
17.1711 | -105.7711
17.1748 | -105.7778
17.1785 | -105.7842
17.1823 | -105.791
17.186 | -105.7965
17.1897 | -105.8048
17.1934 | -105.8115
17.1972 | -105.8162
17.2009 | -105.8274
17.2046 | -105.8334
17.2083 | -105.8438
17.2121 | -105.8515
17.2158 | -105.8573
17.2195 | -105.863
17.2233 | -105.8696
17.227 | -105.8751
17.2307 | -105.8774
17.2344 | -105.878
17.2382 | -105.8792
17.2419 | -105.8793
17.2456 | -105.8822
17.2493 | -105.8846
17.2531 | -105.8838
17.2568 | -105.8826
17.2605 | -105.8812
17.2643 | -105.8771
17.268 | -105.8723
17.2717 | -105.8707
17.2754 | -105.8662
17.2792 | -105.8637
17.2829 | -105.8636
17.2866 | -105.8607
17.2903 | -105.8532
17.2941 | -105.8427
17.2978 | -105.8327
17.3015 | -105.8262
17.3053 | -105.8199
17.309 | -105.8096
17.3127 | -105.8002
17.3164 | -105.7912
17.3202 | -105.7821
17.3239 | -105.7746
17.3276 | -105.7662
17.3313 | -105.758
17.3351 | -105.749
17.3388 | -105.7403
17.3425 | -105.7319
17.3463 | -105.7271
17.35 | -105.7218
17.3537 | -105.7154
17.3574 | -105.7093
17.3612 | -105.7032
17.3649 | -105.6973
17.3686 | -105.6918
17.3723 | -105.686
17.3761 | -105.6818
17.3798 | -105.6778
17.3835 | -105.6737
17.3873 | -105.6715
17.391 | -105.6709
17.3947 | -105.6725
17.3984 | -105.675
17.4022 | -105.6757
17.4059 | -105.6777
17.4096 | -105.6795
17.4133 | -105.6828
17.4171 | -105.6864
17.4208 | -105.6908
17.4245 | -105.6941
17.4283 | -105.6968
17.432 | -105.6992
17.4357 | -105.6973
17.4394 | -105.6997
17.4432 | -105.7058
17.4469 | -105.7059
17.4506 | -105.7048
17.4543 | -105.7027
17.4581 | -105.7001
17.4618 | -105.6967
17.4655 | -105.6931
17.4693 | -105.6872
17.473 | -105.6798
17.4767 | -105.6726
17.4804 | -105.6679
17.4842 | -105.6624
17.4879 | -105.6551
17.4916 | -105.6496
17.4953 | -105.6406
17.4991 | -105.6292
17.5028 | -105.6201
17.5065 | -105.6124
17.5102 | -105.6027
17.514 | -105.5955
17.5177 | -105.5902
17.5214 | -105.5825
17.5252 | -105.5752
17.5289 | -105.5667
17.5326 | -105.5585
17.5363 | -105.5543
17.5401 | -105.5523
17.5438 | -105.5525
17.5475 | -105.5503
17.5512 | -105.5479
17.555 | -105.5488
17.5587 | -105.5529
17.5624 | -105.5558
17.5662 | -105.5585
17.5699 | -105.5634
17.5736 | -105.5698
17.5773 | -105.5784
17.5811 | -105.587
17.5848 | -105.597
17.5885 | -105.6088
17.5922 | -105.6213
17.596 | -105.6317
17.5997 | -105.6454
17.6034 | -105.6616
17.6072 | -105.6767
17.6109 | -105.6911
17.6146 | -105.7071
17.6183 | -105.724
17.6221 | -105.7405
17.6258 | -105.7571
17.6295 | -105.7723
17.6332 | -105.7882
17.637 | -105.8049
17.6407 | -105.8197
17.6444 | -105.8332
17.6482 | -105.8439
17.6519 | -105.8595
17.6556 | -105.8759
17.6593 | -105.891
17.6631 | -105.9038
17.6668 | -105.918
17.6705 | -105.9314
17.6742 | -105.9469
17.678 | -105.9607
17.6817 | -105.9728
17.6854 | -105.9879
17.6892 | -105.998
17.6929 | -106.0061
17.6966 | -106.0179
17.7003 | -106.0292
17.7041 | -106.0399
17.7078 | -106.0487
17.7115 | -106.0593
17.7152 | -106.0698
17.719 | -106.0771
17.7227 | -106.085
17.7264 | -106.0929
17.7302 | -106.0986
17.7339 | -106.1054
17.7376 | -106.1122
17.7413 | -106.1176
17.7451 | -106.1195
17.7488 | -106.1237
17.7525 | -106.1316
17.7562 | -106.1358
17.76 | -106.1366
17.7637 | -106.138
17.7674 | -106.1407
17.7712 | -106.1425
17.7749 | -106.1429
17.7786 | -106.144
17.7823 | -106.1442
17.7861 | -106.1447
17.7898 | -106.1421
17.7935 | -106.1387
17.7972 | -106.1341
17.801 | -106.1285
17.8047 | -106.1217
17.8084 | -106.1122
17.8122 | -106.1011
17.8159 | -106.0911
17.8196 | -106.0823
17.8233 | -106.0725
17.8271 | -106.0605
17.8308 | -106.0473
17.8345 | -106.0331
17.8382 | -106.0206
17.842 | -106.006
17.8457 | -105.99
17.8494 | -105.9746
17.8531 | -105.959
17.8569 | -105.9435
17.8606 | -105.9283
17.8643 | -105.9158
17.8681 | -105.9041
17.8718 | -105.89
17.8755 | -105.8752
17.8792 | -105.8601
17.883 | -105.8484
17.8867 | -105.8376
17.8904 | -105.824
17.8941 | -105.8105
17.8979 | -105.7956
17.9016 | -105.7804
17.9053 | -105.7653
17.9091 | -105.7512
17.9128 | -105.7392
17.9165 | -105.7284
17.9202 | -105.7154
17.924 | -105.7035
17.9277 | -105.6917
17.9314 | -105.6808
17.9351 | -105.6712
17.9389 | -105.6613
17.9426 | -105.6539
17.9463 | -105.6451
17.9501 | -105.6363
17.9538 | -105.6311
17.9575 | -105.6242
17.9612 | -105.6167
17.965 | -105.6124
17.9687 | -105.6105
17.9724 | -105.6097
17.9761 | -105.6062
17.9799 | -105.601
17.9836 | -105.5977
17.9873 | -105.5954
17.9911 | -105.593
17.9948 | -105.592
17.9985 | -105.5915
18.0022 | -105.5917
18.006 | -105.5932
18.0097 | -105.5932
18.0134 | -105.592
18.0171 | -105.5926
18.0209 | -105.5961
18.0246 | -105.5975
18.0283 | -105.5969
18.0321 | -105.5971
18.0358 | -105.598
18.0395 | -105.5993
18.0432 | -105.6003
18.047 | -105.6034
18.0507 | -105.6067
18.0544 | -105.6077
18.0581 | -105.6124
18.0619 | -105.6173
18.0656 | -105.6241
18.0693 | -105.6302
18.0731 | -105.6367
18.0768 | -105.644
18.0805 | -105.6514
18.0842 | -105.6574
18.088 | -105.6634
18.0917 | -105.6742
18.0954 | -105.6829
18.0991 | -105.6928
18.1029 | -105.7013
18.1066 | -105.7092
18.1103 | -105.7176
18.1141 | -105.7276
18.1178 | -105.7361
18.1215 | -105.7466
18.1252 | -105.7571
18.129 | -105.7674
18.1327 | -105.7781
18.1364 | -105.7868
18.1401 | -105.7938
18.1439 | -105.8012
18.1476 | -105.81
18.1513 | -105.8187
18.155 | -105.8276
18.1588 | -105.8345
18.1625 | -105.8376
18.1662 | -105.8372
18.17 | -105.8321
18.1737 | -105.8272
18.1774 | -105.8242
18.1811 | -105.8207
18.1849 | -105.816
18.1886 | -105.8106
18.1923 | -105.8055
18.196 | -105.8008
18.1998 | -105.793
18.2035 | -105.7854
18.2072 | -105.7778
18.211 | -105.769
18.2147 | -105.7618
18.2184 | -105.7534
18.2221 | -105.7417
18.2259 | -105.7302
18.2296 | -105.7205
18.2333 | -105.7118
18.237 | -105.702
18.2408 | -105.6939
18.2445 | -105.6851
18.2482 | -105.6745
18.252 | -105.6631
18.2557 | -105.654
18.2594 | -105.646
18.2631 | -105.6366
18.2669 | -105.6265
18.2706 | -105.6185
18.2743 | -105.6093
18.278 | -105.6004
18.2818 | -105.5956
18.2855 | -105.5889
18.2892 | -105.5823
18.293 | -105.577
18.2967 | -105.5713
18.3004 | -105.5637
18.3041 | -105.5544
18.3079 | -105.5452
18.3116 | -105.538
18.3153 | -105.5349
18.319 | -105.5285
18.3228 | -105.5213
18.3265 | -105.5129
18.3302 | -105.5064
18.334 | -105.4995
18.3377 | -105.4911
18.3414 | -105.4835
18.3451 | -105.4771
18.3489 | -105.4726
18.3526 | -105.4659
18.3563 | -105.46
18.36 | -105.4523
18.3638 | -105.4476
18.3675 | -105.4437
18.3712 | -105.4388
18.375 | -105.4321
18.3787 | -105.4278
18.3824 | -105.4222
18.3861 | -105.4172
18.3899 | -105.4142
18.3936 | -105.4104
18.3973 | -105.4068
18.401 | -105.4046
18.4048 | -105.4009
18.4085 | -105.3966
18.4122 | -105.3928
18.416 | -105.3911
18.4197 | -105.3904
18.4234 | -105.3882
18.4271 | -105.3866
18.4309 | -105.3885
18.4346 | -105.3883
18.4383 | -105.3848
18.442 | -105.3792
18.4458 | -105.3746
18.4495 | -105.3684
18.4532 | -105.3606
18.457 | -105.3522
18.4607 | -105.3441
18.4644 | -105.338
18.4681 | -105.3323
18.4719 | -105.325
18.4756 | -105.3168
18.4793 | -105.3094
18.483 | -105.3022
18.4868 | -105.293
18.4905 | -105.2835
18.4942 | -105.2745
18.4979 | -105.2677
18.5017 | -105.2607
18.5054 | -105.2552
18.5091 | -105.251
18.5129 | -105.2439
18.5166 | -105.2384
18.5203 | -105.2356
18.524 | -105.2331
18.5278 | -105.2322
18.5315 | -105.2352
18.5352 | -105.2395
18.5389 | -105.2435
18.5427 | -105.25
18.5464 | -105.2567
18.5501 | -105.2659
18.5539 | -105.2795
18.5576 | -105.2942
18.5613 | -105.3102
18.565 | -105.3289
18.5688 | -105.3487
18.5725 | -105.3689
18.5762 | -105.3889
18.5799 | -105.4102
18.5837 | -105.4319
18.5874 | -105.4542
18.5911 | -105.4773
18.5949 | -105.5018
18.5986 | -105.5262
18.6023 | -105.5505
18.606 | -105.5732
18.6098 | -105.5964
18.6135 | -105.616
18.6172 | -105.6333
18.6209 | -105.6487
18.6247 | -105.6622
18.6284 | -105.6746
18.6321 | -105.6843
18.6359 | -105.6935
18.6396 | -105.7022
18.6433 | -105.7075
18.647 | -105.7105
18.6508 | -105.7123
18.6545 | -105.7129
18.6582 | -105.7119
18.6619 | -105.7107
18.6657 | -105.7069
18.6694 | -105.7033
18.6731 | -105.6984
18.6769 | -105.6926
18.6806 | -105.6848
18.6843 | -105.6772
18.688 | -105.6717
18.6918 | -105.6668
18.6955 | -105.6579
18.6992 | -105.6483
18.7029 | -105.6417
18.7067 | -105.6346
18.7104 | -105.628
18.7141 | -105.6234
18.7179 | -105.6187
18.7216 | -105.6134
18.7253 | -105.6088
18.729 | -105.6025
18.7328 | -105.5944
18.7365 | -105.5884
18.7402 | -105.5835
18.7439 | -105.5787
18.7477 | -105.5716
18.7514 | -105.5646
18.7551 | -105.5616
18.7589 | -105.5588
18.7626 | -105.5547
18.7663 | -105.5491
18.77 | -105.5434
18.7738 | -105.5382
18.7775 | -105.5365
18.7812 | -105.5322
18.7849 | -105.5257
18.7887 | -105.5196
18.7924 | -105.5145
18.7961 | -105.5076
18.7999 | -105.5006
18.8036 | -105.4929
18.8073 | -105.4849
18.811 | -105.4768
18.8148 | -105.4668
18.8185 | -105.4497
18.8222 | -105.4394
18.8259 | -105.4291
18.8297 | -105.4187
18.8334 | -105.4086
18.8371 | -105.3999
18.8408 | -105.3905
18.8446 | -105.3816
18.8483 | -105.3722
18.852 | -105.3631
18.8558 | -105.3551
18.8595 | -105.348
18.8632 | -105.3408
18.8669 | -105.3349
18.8707 | -105.3304
18.8744 | -105.3261
18.8781 | -105.3196
18.8818 | -105.3151
18.8856 | -105.3124
18.8893 | -105.3106
18.893 | -105.3129
18.8968 | -105.3124
18.9005 | -105.3123
18.9042 | -105.3136
18.9079 | -105.3155
18.9117 | -105.32
18.9154 | -105.326
18.9191 | -105.3318
18.9228 | -105.337
18.9266 | -105.3447
18.9303 | -105.3551
18.934 | -105.3671
18.9378 | -105.3789
18.9415 | -105.3916
18.9452 | -105.4034
18.9489 | -105.4168
18.9527 | -105.4318
18.9564 | -105.4466
18.9601 | -105.4607
18.9638 | -105.4739
18.9676 | -105.4881
18.9713 | -105.5027
18.975 | -105.5162
18.9788 | -105.5281
18.9825 | -105.5407
18.9862 | -105.5556
18.9899 | -105.57
18.9937 | -105.583
18.9974 | -105.5967
19.0011 | -105.6105
19.0048 | -105.6232
19.0086 | -105.6354
19.0123 | -105.6465
19.016 | -105.6593
19.0198 | -105.671
19.0235 | -105.6787
19.0272 | -105.6868
19.0309 | -105.6938
19.0347 | -105.701
19.0384 | -105.7063
19.0421 | -105.7138
19.0458 | -105.7219
19.0496 | -105.728
19.0533 | -105.7319
19.057 | -105.7344
19.0608 | -105.7356
19.0645 | -105.7346
19.0682 | -105.7337
19.0719 | -105.732
19.0757 | -105.7294
19.0794 | -105.7283
19.0831 | -105.7264
19.0868 | -105.724
19.0906 | -105.7217
19.0943 | -105.7143
19.098 | -105.7047
19.1018 | -105.6962
19.1055 | -105.688
19.1092 | -105.6794
19.1129 | -105.6695
19.1167 | -105.6589
19.1204 | -105.6484
19.1241 | -105.6385
19.1278 | -105.63
19.1316 | -105.6222
19.1353 | -105.6136
19.139 | -105.6052
19.1428 | -105.5967
19.1465 | -105.5917
19.1502 | -105.5916
19.1539 | -105.5907
19.1577 | -105.5896
19.1614 | -105.5903
19.1651 | -105.5884
19.1688 | -105.5883
19.1726 | -105.5899
19.1763 | -105.5916
19.18 | -105.5941
19.1837 | -105.5988
19.1875 | -105.6012
19.1912 | -105.6055
19.1949 | -105.6108
19.1987 | -105.616
19.2024 | -105.6202
19.2061 | -105.6248
19.2098 | -105.6304
19.2136 | -105.6355
19.2173 | -105.6429
19.221 | -105.6484
19.2247 | -105.6525
19.2285 | -105.656
19.2322 | -105.6596
19.2359 | -105.6641
19.2397 | -105.6645
19.2434 | -105.662
19.2471 | -105.6618
19.2508 | -105.6597
19.2546 | -105.6553
19.2583 | -105.6487
19.262 | -105.6401
19.2657 | -105.6297
19.2695 | -105.6195
19.2732 | -105.6077
19.2769 | -105.5934
19.2807 | -105.5778
19.2844 | -105.562
19.2881 | -105.542
19.2918 | -105.5219
19.2956 | -105.502
19.2993 | -105.4813
19.303 | -105.4599
19.3067 | -105.4374
19.3105 | -105.4155
19.3142 | -105.3895
19.3179 | -105.3645
19.3217 | -105.3421
19.3254 | -105.3218
19.3291 | -105.3013
19.3328 | -105.2787
19.3366 | -105.2573
19.3403 | -105.2362
19.344 | -105.2168
19.3477 | -105.1978
19.3515 | -105.1786
19.3552 | -105.1683
19.3589 | -105.152
19.3627 | -105.1379
19.3664 | -105.1258
19.3701 | -105.1143
19.3738 | -105.1035
19.3776 | -105.0924
19.3813 | -105.0841
19.385 | -105.0761
19.3887 | -105.0674
19.3925 | -105.0595
19.3962 | -105.0564
19.3999 | -105.0538
19.4037 | -105.0503
19.4074 | -105.0485
19.4111 | -105.0496
19.4148 | -105.0504
19.4186 | -105.0525
19.4223 | -105.0533
19.426 | -105.0551
19.4297 | -105.0588
19.4335 | -105.0618
19.4372 | -105.0637
19.4409 | -105.0689
19.4447 | -105.0747
19.4484 | -105.0796
19.4521 | -105.0822
19.4558 | -105.0847
19.4596 | -105.0873
19.4633 | -105.0938
19.467 | -105.1
19.4707 | -105.1058
19.4745 | -105.1094
19.4782 | -105.1131
19.4819 | -105.1166
19.4856 | -105.1201
19.4894 | -105.125
19.4931 | -105.1296
19.4968 | -105.1321
19.5006 | -105.1344
19.5043 | -105.1396
19.508 | -105.1443
19.5117 | -105.1478
19.5155 | -105.1548
19.5192 | -105.1624
19.5229 | -105.1677
19.5266 | -105.1717
19.5304 | -105.1746
19.5341 | -105.177
19.5378 | -105.18
19.5416 | -105.1847
19.5453 | -105.1868
19.549 | -105.1881
19.5527 | -105.1896
19.5565 | -105.1882
19.5602 | -105.1869
19.5639 | -105.1871
19.5676 | -105.1873
19.5714 | -105.1869
19.5751 | -105.1832
19.5788 | -105.1794
19.5826 | -105.1761
19.5863 | -105.174
19.59 | -105.1748
19.5937 | -105.1721
19.5975 | -105.1681
19.6012 | -105.1624
19.6049 | -105.1582
19.6086 | -105.155
19.6124 | -105.1503
19.6161 | -105.1446
19.6198 | -105.1401
19.6236 | -105.134
19.6273 | -105.1269
19.631 | -105.1215
19.6347 | -105.1154
19.6385 | -105.1083
19.6422 | -105.0994
19.6459 | -105.0889
19.6496 | -105.0825
19.6534 | -105.0767
19.6571 | -105.0722
19.6608 | -105.0629
19.6646 | -105.0561
19.6683 | -105.052
19.672 | -105.0461
19.6757 | -105.0365
19.6795 | -105.0291
19.6832 | -105.0227
19.6869 | -105.0165
19.6906 | -105.0095
19.6944 | -105.0048
19.6981 | -105.0012
19.7018 | -105.0001
19.7056 | -104.9964
19.7093 | -104.9921
19.713 | -104.9921
19.7167 | -104.9925
19.7205 | -104.9924
19.7242 | -104.9936
19.7279 | -104.9942
19.7316 | -104.9979
19.7354 | -105.0011
19.7391 | -105.0041
19.7428 | -105.0044
19.7466 | -105.0039
19.7503 | -105.0057
19.754 | -105.0087
19.7577 | -105.0119
19.7615 | -105.0139
19.7652 | -105.0154
19.7689 | -105.0189
19.7726 | -105.0225
19.7764 | -105.0263
19.7801 | -105.0298
19.7838 | -105.03
19.7876 | -105.032
19.7913 | -105.0359
19.795 | -105.0379
19.7987 | -105.0387
19.8025 | -105.0399
19.8062 | -105.0423
19.8099 | -105.0449
19.8136 | -105.0475
19.8174 | -105.0496
19.8211 | -105.0494
19.8248 | -105.0483
19.8285 | -105.0494
19.8323 | -105.0519
19.836 | -105.0573
19.8397 | -105.0611
19.8435 | -105.0622
19.8472 | -105.0632
19.8509 | -105.0645
19.8546 | -105.0689
19.8584 | -105.073
19.8621 | -105.0769
19.8658 | -105.0796
19.8695 | -105.0804
19.8733 | -105.0831
19.877 | -105.0881
19.8807 | -105.0944
19.8845 | -105.0988
19.8882 | -105.1042
19.8919 | -105.114
19.8956 | -105.1236
19.8994 | -105.1337
19.9031 | -105.1428
19.9068 | -105.1537
19.9105 | -105.1636
19.9143 | -105.1725
19.918 | -105.1844
19.9217 | -105.1966
19.9255 | -105.209
19.9292 | -105.2226
19.9329 | -105.2345
19.9366 | -105.2473
19.9404 | -105.2588
19.9441 | -105.2729
19.9478 | -105.2857
19.9515 | -105.2982
19.9553 | -105.3106
19.959 | -105.3234
19.9627 | -105.3357
19.9665 | -105.3462
19.9702 | -105.3543
19.9739 | -105.3636
19.9776 | -105.3731
19.9814 | -105.3805
19.9851 | -105.3872
19.9888 | -105.3964
19.9925 | -105.4055
19.9963 | -105.4111
20 | -105.416
20.0037 | -105.4193
20.0075 | -105.4221
20.0112 | -105.4252
20.0149 | -105.4253
20.0186 | -105.425
20.0224 | -105.4221
20.0261 | -105.421
20.0298 | -105.4193
20.0335 | -105.4144
20.0373 | -105.4107
20.041 | -105.4039
20.0447 | -105.3969
20.0485 | -105.3902
20.0522 | -105.3811
20.0559 | -105.3708
20.0596 | -105.3635
20.0634 | -105.3555
20.0671 | -105.3485
20.0708 | -105.3392
20.0745 | -105.3297
20.0783 | -105.3202
20.082 | -105.3104
20.0857 | -105.3015
20.0895 | -105.2931
20.0932 | -105.2853
20.0969 | -105.2764
20.1006 | -105.2682
20.1044 | -105.2616
20.1081 | -105.2553
20.1118 | -105.248
20.1155 | -105.24
20.1193 | -105.2333
20.123 | -105.2278
20.1267 | -105.2241
20.1305 | -105.2176
20.1342 | -105.2132
20.1379 | -105.2098
20.1416 | -105.2087
20.1454 | -105.2082
20.1491 | -105.2089
20.1528 | -105.2127
20.1565 | -105.2162
20.1603 | -105.2201
20.164 | -105.2249
20.1677 | -105.2305
20.1715 | -105.236
20.1752 | -105.2425
20.1789 | -105.249
20.1826 | -105.2553
20.1864 | -105.2601
20.1901 | -105.2662
20.1938 | -105.2729
20.1975 | -105.2818
20.2013 | -105.2893
20.205 | -105.2944
20.2087 | -105.3007
20.2124 | -105.3057
20.2162 | -105.3078
20.2199 | -105.31
20.2236 | -105.3145
20.2274 | -105.3171
20.2311 | -105.3196
20.2348 | -105.3203
20.2385 | -105.32
20.2423 | -105.3189
20.246 | -105.3165
20.2497 | -105.3139
20.2534 | -105.3108
20.2572 | -105.3084
20.2609 | -105.3048
20.2646 | -105.2986
20.2684 | -105.292
20.2721 | -105.2831
20.2758 | -105.2732
20.2795 | -105.2635
20.2833 | -105.2535
20.287 | -105.244
20.2907 | -105.2358
20.2944 | -105.2267
20.2982 | -105.2175
20.3019 | -105.2095
20.3056 | -105.2009
20.3094 | -105.1933
20.3131 | -105.1867
20.3168 | -105.1806
20.3205 | -105.1752
20.3243 | -105.1667
20.328 | -105.1596
20.3317 | -105.1537
20.3354 | -105.1463
20.3392 | -105.1396
20.3429 | -105.134
20.3466 | -105.1278
20.3504 | -105.122
20.3541 | -105.1175
20.3578 | -105.1136
20.3615 | -105.1095
20.3653 | -105.1055
20.369 | -105.1041
20.3727 | -105.1024
20.3764 | -105.0998
20.3802 | -105.0965
20.3839 | -105.0956
20.3876 | -105.0974
20.3914 | -105.0995
20.3951 | -105.1018
20.3988 | -105.1029
20.4025 | -105.1034
20.4063 | -105.1042
20.41 | -105.106
20.4137 | -105.1076
20.4174 | -105.1098
20.4212 | -105.1091
20.4249 | -105.1097
20.4286 | -105.109
20.4324 | -105.1085
20.4361 | -105.1089
20.4398 | -105.1088
20.4435 | -105.1064
20.4473 | -105.1051
20.451 | -105.1027
20.4547 | -105.0989
20.4584 | -105.0932
20.4622 | -105.0902
20.4659 | -105.0859
20.4696 | -105.0797
20.4734 | -105.0734
20.4771 | -105.0699
20.4808 | -105.0637
20.4845 | -105.0565
20.4883 | -105.0498
20.492 | -105.0407
20.4957 | -105.0332
20.4994 | -105.0255
20.5032 | -105.0196
20.5069 | -105.0128
20.5106 | -105.0049
20.5144 | -104.9956
20.5181 | -104.9878
20.5218 | -104.9791
20.5255 | -104.9693
20.5293 | -104.9612
20.533 | -104.9549
20.5367 | -104.9481
20.5404 | -104.9408
20.5442 | -104.9356
20.5479 | -104.9294
20.5516 | -104.9218
20.5553 | -104.9158
20.5591 | -104.9106
20.5628 | -104.9073
20.5665 | -104.9025
20.5703 | -104.8977
20.574 | -104.8951
20.5777 | -104.892
20.5814 | -104.888
20.5852 | -104.8845
20.5889 | -104.8818
20.5926 | -104.8786
20.5963 | -104.877
20.6001 | -104.8755
20.6038 | -104.8728
20.6075 | -104.8707
20.6113 | -104.8689
20.615 | -104.8697
20.6187 | -104.8691
20.6224 | -104.867
20.6262 | -104.8679
20.6299 | -104.8689
20.6336 | -104.8692
20.6373 | -104.8688
20.6411 | -104.8698
20.6448 | -104.8709
20.6485 | -104.8723
20.6523 | -104.8743
20.656 | -104.8771
20.6597 | -104.8799
20.6634 | -104.8799
20.6672 | -104.8794
20.6709 | -104.8805
20.6746 | -104.8809
20.6783 | -104.8791
20.6821 | -104.8788
20.6858 | -104.878
20.6895 | -104.8773
20.6933 | -104.8782
20.697 | -104.8754
20.7007 | -104.8716
20.7044 | -104.8694
20.7082 | -104.8676
20.7119 | -104.8659
20.7156 | -104.8613
20.7193 | -104.8571
20.7231 | -104.8524
20.7268 | -104.8467
20.7305 | -104.8405
20.7343 | -104.8354
20.738 | -104.8301
20.7417 | -104.8246
20.7454 | -104.8204
20.7492 | -104.8167
20.7529 | -104.8148
20.7566 | -104.8109
20.7603 | -104.8073
20.7641 | -104.8046
20.7678 | -104.7988
20.7715 | -104.7929
20.7753 | -104.7886
20.779 | -104.7847
20.7827 | -104.7796
20.7864 | -104.7783
20.7902 | -104.7755
20.7939 | -104.773
20.7976 | -104.7703
20.8013 | -104.7704
20.8051 | -104.7695
20.8088 | -104.7688
20.8125 | -104.7669
20.8163 | -104.7659
20.82 | -104.769
20.8237 | -104.7702
20.8274 | -104.7703
20.8312 | -104.771
20.8349 | -104.7722
20.8386 | -104.7741
20.8423 | -104.7772
20.8461 | -104.7817
20.8498 | -104.7861
20.8535 | -104.7914
20.8572 | -104.7963
20.861 | -104.8023
20.8647 | -104.808
20.8684 | -104.8122
20.8722 | -104.8171
20.8759 | -104.8219
20.8796 | -104.8277
20.8833 | -104.8356
20.8871 | -104.8436
20.8908 | -104.8506
20.8945 | -104.8577
20.8982 | -104.8659
20.902 | -104.8738
20.9057 | -104.882
20.9094 | -104.8892
20.9132 | -104.8979
20.9169 | -104.9066
20.9206 | -104.9139
20.9243 | -104.921
20.9281 | -104.9288
20.9318 | -104.9372
20.9355 | -104.9443
20.9392 | -104.9507
20.943 | -104.9562
20.9467 | -104.9617
20.9504 | -104.9674
20.9542 | -104.9704
20.9579 | -104.9762
20.9616 | -104.984
20.9653 | -104.9876
20.9691 | -104.9862
20.9728 | -104.9894
20.9765 | -104.9902
20.9802 | -104.9898
20.984 | -104.9906
20.9877 | -104.9919
20.9914 | -104.9911
20.9952 | -104.9915
20.9989 | -104.9908
21.0026 | -104.99
21.0063 | -104.9886
21.0101 | -104.985
21.0138 | -104.979
21.0175 | -104.9735
21.0212 | -104.9692
21.025 | -104.9622
21.0287 | -104.9568
21.0324 | -104.9522
21.0362 | -104.9479
21.0399 | -104.9401
21.0436 | -104.9316
21.0473 | -104.9238
21.0511 | -104.915
21.0548 | -104.9067
21.0585 | -104.8973
21.0622 | -104.8875
21.066 | -104.8797
21.0697 | -104.8695
21.0734 | -104.8589
21.0772 | -104.8495
21.0809 | -104.8401
21.0846 | -104.8282
21.0883 | -104.8181
21.0921 | -104.8061
21.0958 | -104.7944
21.0995 | -104.7823
21.1032 | -104.7722
21.107 | -104.7611
21.1107 | -104.7528
21.1144 | -104.7444
21.1182 | -104.733
21.1219 | -104.7233
21.1256 | -104.7147
21.1293 | -104.707
21.1331 | -104.6996
21.1368 | -104.6918
21.1405 | -104.6849
21.1442 | -104.6784
21.148 | -104.673
21.1517 | -104.6674
21.1554 | -104.6641
21.1592 | -104.6627
21.1629 | -104.6606
21.1666 | -104.6578
21.1703 | -104.6566
21.1741 | -104.6576
21.1778 | -104.6589
21.1815 | -104.6594
21.1852 | -104.6616
21.189 | -104.6659
21.1927 | -104.6709
21.1964 | -104.6751
21.2001 | -104.6795
21.2039 | -104.6854
21.2076 | -104.6929
21.2113 | -104.701
21.2151 | -104.7086
21.2188 | -104.717
21.2225 | -104.7249
21.2262 | -104.7321
21.23 | -104.7404
21.2337 | -104.7507
21.2374 | -104.7605
21.2411 | -104.7711
21.2449 | -104.7828
21.2486 | -104.794
21.2523 | -104.8057
21.2561 | -104.8164
21.2598 | -104.8268
21.2635 | -104.8394
21.2672 | -104.8519
21.271 | -104.8627
21.2747 | -104.8727
21.2784 | -104.8853
21.2821 | -104.8969
21.2859 | -104.9061
21.2896 | -104.9157
21.2933 | -104.9265
21.2971 | -104.9373
21.3008 | -104.9487
21.3045 | -104.9581
21.3082 | -104.9663
21.312 | -104.9755
21.3157 | -104.986
21.3194 | -104.9942
21.3231 | -105.0011
21.3269 | -105.0085
21.3306 | -105.014
21.3343 | -105.017
21.3381 | -105.022
21.3418 | -105.0281
21.3455 | -105.031
21.3492 | -105.0335
21.353 | -105.0381
21.3567 | -105.0393
21.3604 | -105.0416
21.3641 | -105.0433
21.3679 | -105.043
21.3716 | -105.0447
21.3753 | -105.0461
21.3791 | -105.0453
21.3828 | -105.045
21.3865 | -105.0437
21.3902 | -105.0395
21.394 | -105.0362
21.3977 | -105.0321
21.4014 | -105.0271
21.4051 | -105.0207
21.4089 | -105.0162
21.4126 | -105.0097
21.4163 | -105.0014
21.4201 | -104.9918
21.4238 | -104.9821
21.4275 | -104.9698
21.4312 | -104.9548
21.435 | -104.9401
21.4387 | -104.9252
21.4424 | -104.9087
21.4461 | -104.8903
21.4499 | -104.8698
21.4536 | -104.8473
21.4573 | -104.822
21.4611 | -104.7963
21.4648 | -104.7715
21.4685 | -104.7456
21.4722 | -104.7156
21.476 | -104.682
21.4797 | -104.6495
21.4834 | -104.6173
21.4871 | -104.5846
21.4909 | -104.551
21.4946 | -104.5159
21.4983 | -104.4809
21.5021 | -104.4452
21.5058 | -104.4039
21.5095 | -104.3692
21.5132 | -104.3322
21.517 | -104.2943
21.5207 | -104.2609
21.5244 | -104.2264
21.5281 | -104.1922
21.5319 | -104.1599
21.5356 | -104.1268
21.5393 | -104.0945
21.543 | -104.0653
21.5468 | -104.0371
21.5505 | -104.0097
21.5542 | -103.9834
21.558 | -103.9602
21.5617 | -103.9401
21.5654 | -103.9221
21.5691 | -103.9061
21.5729 | -103.8926
21.5766 | -103.8819
21.5803 | -103.8735
21.584 | -103.8664
21.5878 | -103.8618
21.5915 | -103.8615
21.5952 | -103.8639
21.599 | -103.8683
21.6027 | -103.8745
21.6064 | -103.8829
21.6101 | -103.8948
21.6139 | -103.9103
21.6176 | -103.9285
21.6213 | -103.9489
21.625 | -103.9702
21.6288 | -103.9929
21.6325 | -104.0179
21.6362 | -104.0445
21.64 | -104.0711
21.6437 | -104.0967
21.6474 | -104.1265
21.6511 | -104.1593
21.6549 | -104.1922
21.6586 | -104.2235
21.6623 | -104.2569
21.666 | -104.2893
21.6698 | -104.322
21.6735 | -104.3546
21.6772 | -104.3866
21.681 | -104.4178
21.6847 | -104.4473
21.6884 | -104.4765
21.6921 | -104.5051
21.6959 | -104.5315
21.6996 | -104.5572
21.7033 | -104.5822
21.707 | -104.6039
21.7108 | -104.6239
21.7145 | -104.6421
21.7182 | -104.6601
21.722 | -104.6777
21.7257 | -104.6943
21.7294 | -104.708
21.7331 | -104.7209
21.7369 | -104.7332
21.7406 | -104.7431
21.7443 | -104.7506
21.748 | -104.7561
21.7518 | -104.7603
21.7555 | -104.7652
21.7592 | -104.7703
21.763 | -104.774
21.7667 | -104.7742
21.7704 | -104.7729
21.7741 | -104.7722
21.7779 | -104.7708
21.7816 | -104.7704
21.7853 | -104.7663
21.789 | -104.7601
21.7928 | -104.7542
21.7965 | -104.7487
21.8002 | -104.7435
21.804 | -104.7385
21.8077 | -104.7352
21.8114 | -104.732
21.8151 | -104.7279
21.8189 | -104.7241
21.8226 | -104.7234
21.8263 | -104.7237
21.83 | -104.7207
21.8338 | -104.7185
21.8375 | -104.7178
21.8412 | -104.7168
21.845 | -104.7175
21.8487 | -104.7154
21.8524 | -104.7125
21.8561 | -104.7119
21.8599 | -104.7105
21.8636 | -104.7095
21.8673 | -104.708
21.871 | -104.7032
21.8748 | -104.6988
21.8785 | -104.6947
21.8822 | -104.6897
21.8859 | -104.6831
21.8897 | -104.6772
21.8934 | -104.6729
21.8971 | -104.669
21.9009 | -104.6627
21.9046 | -104.6556
21.9083 | -104.6491
21.912 | -104.6439
21.9158 | -104.6389
21.9195 | -104.6312
21.9232 | -104.6229
21.9269 | -104.6138
21.9307 | -104.6063
21.9344 | -104.5997
21.9381 | -104.5932
21.9419 | -104.5861
21.9456 | -104.5805
21.9493 | -104.5763
21.953 | -104.5737
21.9568 | -104.5682
21.9605 | -104.5607
21.9642 | -104.5546
21.9679 | -104.5501
21.9717 | -104.5449
21.9754 | -104.5396
21.9791 | -104.5351
21.9829 | -104.5341
21.9866 | -104.5322
21.9903 | -104.5292
21.994 | -104.5269
21.9978 | -104.5267
22.0015 | -104.5275
22.0052 | -104.5287
22.0089 | -104.5294
22.0127 | -104.5304
22.0164 | -104.5319
22.0201 | -104.5312
22.0239 | -104.5299
22.0276 | -104.5307
22.0313 | -104.5307
22.035 | -104.5316
22.0388 | -104.5333
22.0425 | -104.5342
22.0462 | -104.532
22.0499 | -104.5326
22.0537 | -104.5318
22.0574 | -104.532
22.0611 | -104.5336
22.0649 | -104.5353
22.0686 | -104.5363
22.0723 | -104.5361
22.076 | -104.5358
22.0798 | -104.5372
22.0835 | -104.5378
22.0872 | -104.5372
22.0909 | -104.538
22.0947 | -104.5375
22.0984 | -104.5369
22.1021 | -104.5366
22.1059 | -104.5357
22.1096 | -104.5349
22.1133 | -104.5331
22.117 | -104.5309
22.1208 | -104.5296
22.1245 | -104.5286
22.1282 | -104.5283
22.1319 | -104.5273
22.1357 | -104.5263
22.1394 | -104.5259
22.1431 | -104.5255
22.1469 | -104.5241
22.1506 | -104.5229
22.1543 | -104.5222
22.158 | -104.5212
22.1618 | -104.5234
22.1655 | -104.5245
22.1692 | -104.5232
22.1729 | -104.5248
22.1767 | -104.5269
22.1804 | -104.5279
22.1841 | -104.5313
22.1878 | -104.5334
22.1916 | -104.5357
22.1953 | -104.5385
22.199 | -104.5418
22.2028 | -104.5434
22.2065 | -104.5426
22.2102 | -104.5439
22.2139 | -104.5451
22.2177 | -104.544
22.2214 | -104.5433
22.2251 | -104.5441
22.2288 | -104.5449
22.2326 | -104.5459
22.2363 | -104.5469
22.24 | -104.5472
22.2438 | -104.548
22.2475 | -104.5488
22.2512 | -104.5482
22.2549 | -104.5484
22.2587 | -104.5482
22.2624 | -104.5481
22.2661 | -104.5467
22.2698 | -104.5434
22.2736 | -104.5423
22.2773 | -104.541
22.281 | -104.5382
22.2848 | -104.537
22.2885 | -104.5366
22.2922 | -104.5346
22.2959 | -104.5321
22.2997 | -104.5309
22.3034 | -104.5305
22.3071 | -104.5305
22.3108 | -104.5269
22.3146 | -104.5228
22.3183 | -104.5214
22.322 | -104.5198
22.3258 | -104.517
22.3295 | -104.5165
22.3332 | -104.5167
22.3369 | -104.514
22.3407 | -104.5108
22.3444 | -104.5096
22.3481 | -104.5076
22.3518 | -104.5047
22.3556 | -104.504
22.3593 | -104.504
22.363 | -104.5035
22.3668 | -104.5017
22.3705 | -104.5002
22.3742 | -104.4979
22.3779 | -104.4957
22.3817 | -104.4938
22.3854 | -104.4926
22.3891 | -104.4937
22.3928 | -104.493
22.3966 | -104.4899
22.4003 | -104.4899
22.404 | -104.4904
22.4078 | -104.4914
22.4115 | -104.4942
22.4152 | -104.4982
22.4189 | -104.5034
22.4227 | -104.5088
22.4264 | -104.5142
22.4301 | -104.5211
22.4338 | -104.5282
22.4376 | -104.5355
22.4413 | -104.543
22.445 | -104.5509
22.4488 | -104.5588
22.4525 | -104.5659
22.4562 | -104.5741
22.4599 | -104.5835
22.4637 | -104.5943
22.4674 | -104.6045
22.4711 | -104.6154
22.4748 | -104.6268
22.4786 | -104.6355
22.4823 | -104.6465
22.486 | -104.6576
22.4898 | -104.6654
22.4935 | -104.6681
22.4972 | -104.6691
22.5009 | -104.6669
22.5047 | -104.6656
22.5084 | -104.6501
22.5121 | -104.6273
22.5158 | -104.5988
22.5196 | -104.5691
22.5233 | -104.5352
22.527 | -104.5022
22.5307 | -104.4629
22.5345 | -104.414
22.5382 | -104.3662
22.5419 | -104.3217
22.5457 | -104.2727
22.5494 | -104.2226
22.5531 | -104.1681
22.5568 | -104.1097
22.5606 | -104.0466
22.5643 | -103.981
22.568 | -103.9137
22.5717 | -103.8426
22.5755 | -103.7676
22.5792 | -103.6901
22.5829 | -103.5878
22.5867 | -103.5087
22.5904 | -103.4303
22.5941 | -103.3517
22.5978 | -103.2723
22.6016 | -103.1935
22.6053 | -103.114
22.609 | -103.0354
22.6127 | -102.9586
22.6165 | -102.8833
22.6202 | -102.8093
22.6239 | -102.7377
22.6277 | -102.6685
22.6314 | -102.6051
22.6351 | -102.5454
22.6388 | -102.4874
22.6426 | -102.4333
22.6463 | -102.3828
22.65 | -102.3367
22.6537 | -102.2968
22.6575 | -102.2628
22.6612 | -102.2328
22.6649 | -102.2069
22.6687 | -102.186
22.6724 | -102.1707
22.6761 | -102.1607
22.6798 | -102.1563
22.6836 | -102.1568
22.6873 | -102.162
22.691 | -102.1735
22.6947 | -102.1897
22.6985 | -102.2112
22.7022 | -102.2387
22.7059 | -102.2695
22.7097 | -102.3044
22.7134 | -102.3455
22.7171 | -102.3919
22.7208 | -102.4407
22.7246 | -102.4931
22.7283 | -102.5481
22.732 | -102.6047
22.7357 | -102.6635
22.7395 | -102.7247
22.7432 | -102.7889
22.7469 | -102.856
22.7507 | -102.9248
22.7544 | -102.9945
22.7581 | -103.0631
22.7618 | -103.1328
22.7656 | -103.203
22.7693 | -103.2731
22.773 | -103.3427
22.7767 | -103.4112
22.7805 | -103.4785
22.7842 | -103.5453
22.7879 | -103.6116
22.7917 | -103.6745
22.7954 | -103.7361
22.7991 | -103.7958
22.8028 | -103.8503
22.8066 | -103.9045
22.8103 | -103.9577
22.814 | -104.0123
22.8177 | -104.0653
22.8215 | -104.116
22.8252 | -104.1606
22.8289 | -104.2111
22.8327 | -104.262
22.8364 | -104.3103
22.8401 | -104.3526
22.8438 | -104.39
22.8476 | -104.4175
22.8513 | -104.4447
22.855 | -104.4753
22.8587 | -104.4976
22.8625 | -104.5074
22.8662 | -104.517
22.8699 | -104.525
22.8736 | -104.5332
22.8774 | -104.5419
22.8811 | -104.5489
22.8848 | -104.5552
22.8886 | -104.5605
22.8923 | -104.566
22.896 | -104.5703
22.8997 | -104.5748
22.9035 | -104.5771
22.9072 | -104.5783
22.9109 | -104.5812
22.9146 | -104.5832
22.9184 | -104.5815
22.9221 | -104.5772
22.9258 | -104.5722
22.9296 | -104.5683
22.9333 | -104.5651
22.937 | -104.5607
22.9407 | -104.5544
22.9445 | -104.5467
22.9482 | -104.5394
22.9519 | -104.5336
22.9556 | -104.524
22.9594 | -104.5128
22.9631 | -104.5024
22.9668 | -104.4912
22.9706 | -104.4815
22.9743 | -104.4707
22.978 | -104.4575
22.9817 | -104.4431
22.9855 | -104.4277
22.9892 | -104.4153
22.9929 | -104.4044
22.9966 | -104.3908
23.0004 | -104.3777
23.0041 | -104.3622
23.0078 | -104.3463
23.0116 | -104.3294
23.0153 | -104.3092
23.019 | -104.2933
23.0227 | -104.2767
23.0265 | -104.2568
23.0302 | -104.2376
23.0339 | -104.2196
23.0376 | -104.2009
23.0414 | -104.1813
23.0451 | -104.1611
23.0488 | -104.141
23.0526 | -104.1215
23.0563 | -104.102
23.06 | -104.0847
23.0637 | -104.0661
23.0675 | -104.0463
23.0712 | -104.0238
23.0749 | -104.0033
23.0786 | -103.983
23.0824 | -103.9633
23.0861 | -103.9451
23.0898 | -103.9266
23.0936 | -103.9077
23.0973 | -103.8881
23.101 | -103.8691
23.1047 | -103.852
23.1085 | -103.8348
23.1122 | -103.818
23.1159 | -103.803
23.1196 | -103.7947
23.1234 | -103.7813
23.1271 | -103.7702
23.1308 | -103.7598
23.1346 | -103.7486
23.1383 | -103.7391
23.142 | -103.7314
23.1457 | -103.7239
23.1495 | -103.7167
23.1532 | -103.7123
23.1569 | -103.7094
23.1606 | -103.7076
23.1644 | -103.706
23.1681 | -103.7046
23.1718 | -103.705
23.1756 | -103.7063
23.1793 | -103.711
23.183 | -103.7156
23.1867 | -103.719
23.1905 | -103.7238
23.1942 | -103.7299
23.1979 | -103.7358
23.2016 | -103.7418
23.2054 | -103.7472
23.2091 | -103.7537
23.2128 | -103.7609
23.2165 | -103.7688
23.2203 | -103.7768
23.224 | -103.7875
23.2277 | -103.7963
23.2315 | -103.806
23.2352 | -103.8162
23.2389 | -103.824
23.2426 | -103.8326
23.2464 | -103.8425
23.2501 | -103.8525
23.2538 | -103.8605
23.2575 | -103.8669
23.2613 | -103.871
23.265 | -103.8757
23.2687 | -103.8802
23.2725 | -103.8864
23.2762 | -103.8928
23.2799 | -103.8967
23.2836 | -103.9017
23.2874 | -103.9059
23.2911 | -103.9078
23.2948 | -103.9103
23.2985 | -103.9107
23.3023 | -103.9098
23.306 | -103.9096
23.3097 | -103.9098
23.3135 | -103.9104
23.3172 | -103.9092
23.3209 | -103.908
23.3246 | -103.9094
23.3284 | -103.9085
23.3321 | -103.909
23.3358 | -103.9096
23.3395 | -103.9113
23.3433 | -103.9136
23.347 | -103.9132
23.3507 | -103.9138
23.3545 | -103.9159
23.3582 | -103.9176
23.3619 | -103.9179
23.3656 | -103.9191
23.3694 | -103.9216
23.3731 | -103.9237
23.3768 | -103.9262
23.3805 | -103.9292
23.3843 | -103.9345
23.388 | -103.9377
23.3917 | -103.9427
23.3955 | -103.9483
23.3992 | -103.955
23.4029 | -103.9623
23.4066 | -103.9714
23.4104 | -103.9807
23.4141 | -103.9875
23.4178 | -103.9942
23.4215 | -104.0018
23.4253 | -104.0097
23.429 | -104.0178
23.4327 | -104.0241
23.4365 | -104.0302
23.4402 | -104.0371
23.4439 | -104.0453
23.4476 | -104.0533
23.4514 | -104.0604
23.4551 | -104.0662
23.4588 | -104.0722
23.4625 | -104.0796
23.4663 | -104.0868
23.47 | -104.0937
23.4737 | -104.0994
23.4775 | -104.1055
23.4812 | -104.1111
23.4849 | -104.1153
23.4886 | -104.1208
23.4924 | -104.1246
23.4961 | -104.127
23.4998 | -104.1311
23.5035 | -104.1344
23.5073 | -104.1355
23.511 | -104.1348
23.5147 | -104.1349
23.5184 | -104.1357
23.5222 | -104.1376
23.5259 | -104.1383
23.5296 | -104.1398
23.5334 | -104.1404
23.5371 | -104.142
23.5408 | -104.1436
23.5445 | -104.1442
23.5483 | -104.1445
23.552 | -104.1445
23.5557 | -104.1431
23.5594 | -104.1436
23.5632 | -104.1446
23.5669 | -104.1443
23.5706 | -104.1453
23.5744 | -104.1459
23.5781 | -104.1457
23.5818 | -104.1472
23.5855 | -104.1523
23.5893 | -104.1584
23.593 | -104.1653
23.5967 | -104.1707
23.6004 | -104.1766
23.6042 | -104.1806
23.6079 | -104.1842
23.6116 | -104.1901
23.6154 | -104.1962
23.6191 | -104.2047
23.6228 | -104.2146
23.6265 | -104.2239
23.6303 | -104.2345
23.634 | -104.2446
23.6377 | -104.2532
23.6414 | -104.263
23.6452 | -104.2732
23.6489 | -104.2835
23.6526 | -104.2943
23.6564 | -104.3102
23.6601 | -104.3211
23.6638 | -104.3334
23.6675 | -104.3465
23.6713 | -104.3609
23.675 | -104.3749
23.6787 | -104.3889
23.6824 | -104.4037
23.6862 | -104.4174
23.6899 | -104.4312
23.6936 | -104.4465
23.6974 | -104.4612
23.7011 | -104.4748
23.7048 | -104.4869
23.7085 | -104.5023
23.7123 | -104.5177
23.716 | -104.5325
23.7197 | -104.5445
23.7234 | -104.5572
23.7272 | -104.568
23.7309 | -104.5761
23.7346 | -104.5851
23.7384 | -104.5939
23.7421 | -104.6021
23.7458 | -104.6117
23.7495 | -104.6217
23.7533 | -104.6306
23.757 | -104.6411
23.7607 | -104.6508
23.7644 | -104.656
23.7682 | -104.6616
23.7719 | -104.6667
23.7756 | -104.6691
23.7794 | -104.668
23.7831 | -104.6652
23.7868 | -104.6646
23.7905 | -104.6637
23.7943 | -104.6601
23.798 | -104.6573
23.8017 | -104.6533
23.8054 | -104.645
23.8092 | -104.636
23.8129 | -104.6277
23.8166 | -104.6216
23.8204 | -104.6134
23.8241 | -104.6012
23.8278 | -104.5904
23.8315 | -104.5789
23.8353 | -104.5643
23.839 | -104.5529
23.8427 | -104.5426
23.8464 | -104.5315
23.8502 | -104.5193
23.8539 | -104.5065
23.8576 | -104.4934
23.8613 | -104.4795
23.8651 | -104.4655
23.8688 | -104.4532
23.8725 | -104.439
23.8763 | -104.4255
23.88 | -104.4151
23.8837 | -104.4056
23.8874 | -104.3976
23.8912 | -104.3902
23.8949 | -104.3835
23.8986 | -104.3767
23.9023 | -104.3729
23.9061 | -104.369
23.9098 | -104.3645
23.9135 | -104.3615
23.9173 | -104.3614
23.921 | -104.3603
23.9247 | -104.3591
23.9284 | -104.3601
23.9322 | -104.3637
23.9359 | -104.3688
23.9396 | -104.3758
23.9433 | -104.3841
23.9471 | -104.3923
23.9508 | -104.4004
23.9545 | -104.4094
23.9583 | -104.4205
23.962 | -104.4319
23.9657 | -104.4447
23.9694 | -104.4583
23.9732 | -104.4723
23.9769 | -104.4861
23.9806 | -104.5
23.9843 | -104.5141
23.9881 | -104.5287
23.9918 | -104.5435
23.9955 | -104.5585
23.9993 | -104.5745
24.003 | -104.5907
24.0067 | -104.6062
24.0104 | -104.6211
24.0142 | -104.6361
24.0179 | -104.6518
24.0216 | -104.6666
24.0253 | -104.6826
24.0291 | -104.6995
24.0328 | -104.7156
24.0365 | -104.731
24.0403 | -104.7454
24.044 | -104.759
24.0477 | -104.7703
24.0514 | -104.7817
24.0552 | -104.7923
24.0589 | -104.8042
24.0626 | -104.8166
24.0663 | -104.828
24.0701 | -104.8394
24.0738 | -104.8502
24.0775 | -104.8581
24.0813 | -104.8636
24.085 | -104.8678
24.0887 | -104.8724
24.0924 | -104.8769
24.0962 | -104.8776
24.0999 | -104.88
24.1036 | -104.8844
24.1073 | -104.8867
24.1111 | -104.8871
24.1148 | -104.8883
24.1185 | -104.8908
24.1223 | -104.8919
24.126 | -104.8896
24.1297 | -104.888
24.1334 | -104.8865
24.1372 | -104.8829
24.1409 | -104.8784
24.1446 | -104.8748
24.1483 | -104.8704
24.1521 | -104.8639
24.1558 | -104.854
24.1595 | -104.8458
24.1633 | -104.8412
24.167 | -104.8345
24.1707 | -104.8286
24.1744 | -104.8239
24.1782 | -104.816
24.1819 | -104.8077
24.1856 | -104.7997
24.1893 | -104.7926
24.1931 | -104.7845
24.1968 | -104.7806
24.2005 | -104.7725
24.2042 | -104.7664
24.208 | -104.7619
24.2117 | -104.757
24.2154 | -104.7543
24.2192 | -104.7503
24.2229 | -104.744
24.2266 | -104.7376
24.2303 | -104.7319
24.2341 | -104.7269
24.2378 | -104.723
24.2415 | -104.7194
24.2452 | -104.7139
24.249 | -104.7116
24.2527 | -104.7107
24.2564 | -104.709
24.2602 | -104.7048
24.2639 | -104.7012
24.2676 | -104.6984
24.2713 | -104.6944
24.2751 | -104.6856
24.2788 | -104.6741
24.2825 | -104.6543
24.2862 | -104.6213
24.29 | -104.5783
24.2937 | -104.5217
24.2974 | -104.4618
24.3012 | -104.4122
24.3049 | -104.3749
24.3086 | -104.3381
24.3123 | -104.3001
24.3161 | -104.2595
24.3198 | -104.2169
24.3235 | -104.1711
24.3272 | -104.1228
24.331 | -104.0712
24.3347 | -104.0161
24.3384 | -103.9585
24.3422 | -103.9
24.3459 | -103.8389
24.3496 | -103.7752
24.3533 | -103.7108
24.3571 | -103.6451
24.3608 | -103.5772
24.3645 | -103.5072
24.3682 | -103.4353
24.372 | -103.3628
24.3757 | -103.291
24.3794 | -103.2201
24.3832 | -103.1491
24.3869 | -103.078
24.3906 | -103.0073
24.3943 | -102.9368
24.3981 | -102.8683
24.4018 | -102.8015
24.4055 | -102.7364
24.4092 | -102.6749
24.413 | -102.6173
24.4167 | -102.5607
24.4204 | -102.5072
24.4242 | -102.4583
24.4279 | -102.4129
24.4316 | -102.3708
24.4353 | -102.3323
24.4391 | -102.2993
24.4428 | -102.2702
24.4465 | -102.2454
24.4502 | -102.2251
24.454 | -102.2099
24.4577 | -102.1993
24.4614 | -102.1927
24.4652 | -102.1911
24.4689 | -102.194
24.4726 | -102.2011
24.4763 | -102.2134
24.4801 | -102.2299
24.4838 | -102.2498
24.4875 | -102.275
24.4912 | -102.3041
24.495 | -102.3368
24.4987 | -102.374
24.5024 | -102.4136
24.5062 | -102.4563
24.5099 | -102.5021
24.5136 | -102.549
24.5173 | -102.5987
24.5211 | -102.6508
24.5248 | -102.7048
24.5285 | -102.7615
24.5322 | -102.8188
24.536 | -102.8793
24.5397 | -102.938
24.5434 | -102.9958
24.5471 | -103.0549
24.5509 | -103.1138
24.5546 | -103.1693
24.5583 | -103.2226
24.5621 | -103.2719
24.5658 | -103.3163
24.5695 | -103.3507
24.5732 | -103.3711
24.577 | -103.3636
24.5807 | -103.3242
24.5844 | -103.2699
24.5881 | -103.1902
24.5919 | -103.1091
24.5956 | -103.0439
24.5993 | -102.9738
24.6031 | -102.9112
24.6068 | -102.8462
24.6105 | -102.77
24.6142 | -102.683
24.618 | -102.5856
24.6217 | -102.4826
24.6254 | -102.3768
24.6291 | -102.2701
24.6329 | -102.1637
24.6366 | -102.056
24.6403 | -101.9464
24.6441 | -101.8292
24.6478 | -101.6992
24.6515 | -101.5589
24.6552 | -101.4155
24.659 | -101.2717
24.6627 | -101.1281
24.6664 | -100.985
24.6701 | -100.8448
24.6739 | -100.7069
24.6776 | -100.5713
24.6813 | -100.439
24.6851 | -100.3101
24.6888 | -100.1846
24.6925 | -100.0641
24.6962 | -99.9488
24.7 | -99.8373
24.7037 | -99.731
24.7074 | -99.6306
24.7111 | -99.5362
24.7149 | -99.4486
24.7186 | -99.3677
24.7223 | -99.2931
24.7261 | -99.2251
24.7298 | -99.1644
24.7335 | -99.165
24.7372 | -99.1208
24.741 | -99.0838
24.7447 | -99.0538
24.7484 | -99.0305
24.7521 | -99.014
24.7559 | -99.0051
24.7596 | -99.0034
24.7633 | -99.0091
24.7671 | -99.0224
24.7708 | -99.0429
24.7745 | -99.0707
24.7782 | -99.1059
24.782 | -99.1485
24.7857 | -99.198
24.7894 | -99.2542
24.7931 | -99.3168
24.7969 | -99.3865
24.8006 | -99.4628
24.8043 | -99.5452
24.8081 | -99.6335
24.8118 | -99.7287
24.8155 | -99.8299
24.8192 | -99.9369
24.823 | -100.0488
24.8267 | -100.165
24.8304 | -100.2866
24.8341 | -100.4138
24.8379 | -100.544
24.8416 | -100.6775
24.8453 | -100.8145
24.849 | -100.9537
24.8528 | -101.0956
24.8565 | -101.2381
24.8602 | -101.3828
24.864 | -101.5303
24.8677 | -101.6765
24.8714 | -101.8218
24.8751 | -101.9664
24.8789 | -102.1085
24.8826 | -102.2497
24.8863 | -102.3876
24.89 | -102.5235
24.8938 | -102.6572
24.8975 | -102.7866
24.9012 | -102.9114
24.905 | -103.0333
24.9087 | -103.1503
24.9124 | -103.2632
24.9161 | -103.3739
24.9199 | -103.4842
24.9236 | -103.5983
24.9273 | -103.7256
24.931 | -103.8656
24.9348 | -103.9969
24.9385 | -104.1327
24.9422 | -104.2441
24.946 | -104.3092
24.9497 | -104.3613
24.9534 | -104.3884
24.9571 | -104.4043
24.9609 | -104.4217
24.9646 | -104.4394
24.9683 | -104.4611
24.972 | -104.4789
24.9758 | -104.4893
24.9795 | -104.4956
24.9832 | -104.4988
24.987 | -104.4985
24.9907 | -104.4986
24.9944 | -104.4984
24.9981 | -104.4984
25.0019 | -104.4985
25.0056 | -104.4973
25.0093 | -104.4968
25.013 | -104.4989
25.0168 | -104.5013
25.0205 | -104.5015
25.0242 | -104.5008
25.028 | -104.5009
25.0317 | -104.5036
25.0354 | -104.5052
25.0391 | -104.5061
25.0429 | -104.5069
25.0466 | -104.5077
25.0503 | -104.5107
25.054 | -104.5144
25.0578 | -104.5149
25.0615 | -104.5166
25.0652 | -104.5162
25.069 | -104.5148
25.0727 | -104.5124
25.0764 | -104.5128
25.0801 | -104.5137
25.0839 | -104.5131
25.0876 | -104.5144
25.0913 | -104.5169
25.095 | -104.5189
25.0988 | -104.5199
25.1025 | -104.5207
25.1062 | -104.5232
25.11 | -104.5243
25.1137 | -104.5245
25.1174 | -104.5221
25.1211 | -104.5213
25.1249 | -104.5203
25.1286 | -104.5169
25.1323 | -104.5132
25.136 | -104.5108
25.1398 | -104.5107
25.1435 | -104.5111
25.1472 | -104.512
25.151 | -104.5109
25.1547 | -104.5075
25.1584 | -104.5057
25.1621 | -104.5049
25.1659 | -104.5024
25.1696 | -104.5013
25.1733 | -104.4995
25.177 | -104.4985
25.1808 | -104.4971
25.1845 | -104.4945
25.1882 | -104.4916
25.1919 | -104.4904
25.1957 | -104.489
25.1994 | -104.4874
25.2031 | -104.4874
25.2069 | -104.4868
25.2106 | -104.4865
25.2143 | -104.4842
25.218 | -104.4816
25.2218 | -104.4813
25.2255 | -104.4803
25.2292 | -104.4786
25.2329 | -104.4788
25.2367 | -104.4791
25.2404 | -104.4787
25.2441 | -104.4793
25.2479 | -104.4787
25.2516 | -104.4803
25.2553 | -104.4808
25.259 | -104.4803
25.2628 | -104.4819
25.2665 | -104.484
25.2702 | -104.4816
25.2739 | -104.482
25.2777 | -104.4832
25.2814 | -104.4828
25.2851 | -104.4825
25.2889 | -104.4843
25.2926 | -104.4859
25.2963 | -104.4859
25.3 | -104.4851
25.3038 | -104.4862
25.3075 | -104.4854
25.3112 | -104.4836
25.3149 | -104.4847
25.3187 | -104.4838
25.3224 | -104.4828
25.3261 | -104.4809
25.3299 | -104.4814
25.3336 | -104.4796
25.3373 | -104.4791
25.341 | -104.4797
25.3448 | -104.4798
25.3485 | -104.4788
25.3522 | -104.4782
25.3559 | -104.4764
25.3597 | -104.4754
25.3634 | -104.4779
25.3671 | -104.4786
25.3709 | -104.4779
25.3746 | -104.4775
25.3783 | -104.4776
25.382 | -104.4776
25.3858 | -104.4783
25.3895 | -104.4767
25.3932 | -104.4757
25.3969 | -104.4758
25.4007 | -104.4745
25.4044 | -104.4724
25.4081 | -104.4699
25.4119 | -104.466
25.4156 | -104.4656
25.4193 | -104.462
25.423 | -104.4579
25.4268 | -104.4546
25.4305 | -104.4533
25.4342 | -104.4493
25.4379 | -104.4453
25.4417 | -104.4414
25.4454 | -104.4349
25.4491 | -104.4264
25.4529 | -104.4177
25.4566 | -104.4114
25.4603 | -104.4061
25.464 | -104.3993
25.4678 | -104.3933
25.4715 | -104.3873
25.4752 | -104.382
25.4789 | -104.3753
25.4827 | -104.3684
25.4864 | -104.3631
25.4901 | -104.3576
25.4938 | -104.3516
25.4976 | -104.3437
25.5013 | -104.3349
25.505 | -104.3265
25.5088 | -104.3191
25.5125 | -104.313
25.5162 | -104.3063
25.5199 | -104.2994
25.5237 | -104.2904
25.5274 | -104.2822
25.5311 | -104.2754
25.5348 | -104.2652
25.5386 | -104.2535
25.5423 | -104.244
25.546 | -104.2337
25.5498 | -104.2231
25.5535 | -104.2145
25.5572 | -104.2051
25.5609 | -104.1932
25.5647 | -104.1825
25.5684 | -104.1727
25.5721 | -104.1635
25.5758 | -104.1533
25.5796 | -104.1417
25.5833 | -104.1298
25.587 | -104.1195
25.5908 | -104.1091
25.5945 | -104.0983
25.5982 | -104.0889
25.6019 | -104.0805
25.6057 | -104.0722
25.6094 | -104.0643
25.6131 | -104.0567
25.6168 | -104.0494
25.6206 | -104.042
25.6243 | -104.035
25.628 | -104.0308
25.6318 | -104.0273
25.6355 | -104.0228
25.6392 | -104.0194
25.6429 | -104.0165
25.6467 | -104.0139
25.6504 | -104.0114
25.6541 | -104.0094
25.6578 | -104.0084
25.6616 | -104.0101
25.6653 | -104.0095
25.669 | -104.0113
25.6728 | -104.0132
25.6765 | -104.0155
25.6802 | -104.0212
25.6839 | -104.0264
25.6877 | -104.0309
25.6914 | -104.0362
25.6951 | -104.0437
25.6988 | -104.0489
25.7026 | -104.0537
25.7063 | -104.0603
25.71 | -104.067
25.7138 | -104.0736
25.7175 | -104.0821
25.7212 | -104.0893
25.7249 | -104.0952
25.7287 | -104.1022
25.7324 | -104.1098
25.7361 | -104.117
25.7398 | -104.1246
25.7436 | -104.133
25.7473 | -104.1404
25.751 | -104.1478
25.7548 | -104.1559
25.7585 | -104.1632
25.7622 | -104.1708
25.7659 | -104.1783
25.7697 | -104.1857
25.7734 | -104.1931
25.7771 | -104.1998
25.7808 | -104.2067
25.7846 | -104.2131
25.7883 | -104.2177
25.792 | -104.222
25.7958 | -104.2242
25.7995 | -104.2283
25.8032 | -104.2342
25.8069 | -104.2368
25.8107 | -104.2423
25.8144 | -104.2482
25.8181 | -104.2531
25.8218 | -104.2563
25.8256 | -104.2594
25.8293 | -104.2639
25.833 | -104.2672
25.8367 | -104.2709
25.8405 | -104.2745
25.8442 | -104.2779
25.8479 | -104.2823
25.8517 | -104.2849
25.8554 | -104.2868
25.8591 | -104.2891
25.8628 | -104.2937
25.8666 | -104.2987
25.8703 | -104.3024
25.874 | -104.3059
25.8777 | -104.3101
25.8815 | -104.3133
25.8852 | -104.3164
25.8889 | -104.3223
25.8927 | -104.3258
25.8964 | -104.3287
25.9001 | -104.3343
25.9038 | -104.3427
25.9076 | -104.349
25.9113 | -104.355
25.915 | -104.3595
25.9187 | -104.3628
25.9225 | -104.3661
25.9262 | -104.3701
25.9299 | -104.3725
25.9337 | -104.3755
25.9374 | -104.378
25.9411 | -104.3817
25.9448 | -104.388
25.9486 | -104.3946
25.9523 | -104.3991
25.956 | -104.4022
25.9597 | -104.4056
25.9635 | -104.4085
25.9672 | -104.4108
25.9709 | -104.414
25.9747 | -104.4164
25.9784 | -104.4203
25.9821 | -104.4247
25.9858 | -104.4299
25.9896 | -104.434
25.9933 | -104.4361
25.997 | -104.4387
26.0007 | -104.442
26.0045 | -104.4458
26.0082 | -104.4504
26.0119 | -104.4551
26.0157 | -104.4594
26.0194 | -104.4626
26.0231 | -104.4647
26.0268 | -104.4662
26.0306 | -104.4713
26.0343 | -104.4745
26.038 | -104.4771
26.0417 | -104.478
26.0455 | -104.4796
26.0492 | -104.4834
26.0529 | -104.4849
26.0567 | -104.4857
26.0604 | -104.4856
26.0641 | -104.4865
26.0678 | -104.4868
26.0716 | -104.4874
26.0753 | -104.4892
26.079 | -104.4921
26.0827 | -104.4931
26.0865 | -104.4942
26.0902 | -104.4964
26.0939 | -104.4986
26.0977 | -104.4996
26.1014 | -104.4999
26.1051 | -104.4993
26.1088 | -104.4991
26.1126 | -104.4988
26.1163 | -104.4977
26.12 | -104.4955
26.1237 | -104.4941
26.1275 | -104.4934
26.1312 | -104.4931
26.1349 | -104.4915
26.1387 | -104.4904
26.1424 | -104.4891
26.1461 | -104.4879
26.1498 | -104.4863
26.1536 | -104.4847
26.1573 | -104.4816
26.161 | -104.4793
26.1647 | -104.4749
26.1685 | -104.4704
26.1722 | -104.4663
26.1759 | -104.4628
26.1796 | -104.4586
26.1834 | -104.453
26.1871 | -104.4458
26.1908 | -104.44
26.1946 | -104.4347
26.1983 | -104.428
26.202 | -104.4217
26.2057 | -104.4148
26.2095 | -104.4076
26.2132 | -104.401
26.2169 | -104.3927
26.2206 | -104.3866
26.2244 | -104.3806
26.2281 | -104.3739
26.2318 | -104.3669
26.2356 | -104.3596
26.2393 | -104.3539
26.243 | -104.3475
26.2467 | -104.3413
26.2505 | -104.3353
26.2542 | -104.3302
26.2579 | -104.3224
26.2616 | -104.3143
26.2654 | -104.3062
26.2691 | -104.2987
26.2728 | -104.2919
26.2766 | -104.2848
26.2803 | -104.2789
26.284 | -104.2732
26.2877 | -104.2665
26.2915 | -104.261
26.2952 | -104.2561
26.2989 | -104.2518
26.3026 | -104.247
26.3064 | -104.2434
26.3101 | -104.2385
26.3138 | -104.2309
26.3176 | -104.2232
26.3213 | -104.2163
26.325 | -104.2119
26.3287 | -104.2079
26.3325 | -104.203
26.3362 | -104.1967
26.3399 | -104.1924
26.3436 | -104.1874
26.3474 | -104.1819
26.3511 | -104.1769
26.3548 | -104.1723
26.3586 | -104.1671
26.3623 | -104.1629
26.366 | -104.1599
26.3697 | -104.155
26.3735 | -104.1484
26.3772 | -104.1427
26.3809 | -104.1373
26.3846 | -104.1321
26.3884 | -104.1258
26.3921 | -104.1228
26.3958 | -104.1209
26.3996 | -104.1162
26.4033 | -104.1106
26.407 | -104.1053
26.4107 | -104.1004
26.4145 | -104.0953
26.4182 | -104.0898
26.4219 | -104.084
26.4256 | -104.0794
26.4294 | -104.0757
26.4331 | -104.0729
26.4368 | -104.0694
26.4406 | -104.0663
26.4443 | -104.0623
26.448 | -104.058
26.4517 | -104.0536
26.4555 | -104.05
26.4592 | -104.045
26.4629 | -104.0394
26.4666 | -104.0348
26.4704 | -104.0309
26.4741 | -104.0279
26.4778 | -104.0249
26.4816 | -104.0229
26.4853 | -104.0226
26.489 | -104.0235
26.4927 | -104.0233
26.4965 | -104.0229
26.5002 | -104.0216
26.5039 | -104.0211
26.5076 | -104.022
26.5114 | -104.0221
26.5151 | -104.0217
26.5188 | -104.0231
26.5225 | -104.0242
26.5263 | -104.0249
26.53 | -104.0259
26.5337 | -104.0285
26.5375 | -104.0296
26.5412 | -104.0315
26.5449 | -104.0349
26.5486 | -104.0373
26.5524 | -104.0403
26.5561 | -104.0444
26.5598 | -104.0502
26.5635 | -104.0544
26.5673 | -104.0572
26.571 | -104.0611
26.5747 | -104.064
26.5785 | -104.0674
26.5822 | -104.0715
26.5859 | -104.0755
26.5896 | -104.0808
26.5934 | -104.0853
26.5971 | -104.0889
26.6008 | -104.0899
26.6045 | -104.0915
26.6083 | -104.0953
26.612 | -104.0983
26.6157 | -104.1025
26.6195 | -104.1071
26.6232 | -104.1119
26.6269 | -104.115
26.6306 | -104.1195
26.6344 | -104.1245
26.6381 | -104.1289
26.6418 | -104.1334
26.6455 | -104.137
26.6493 | -104.1398
26.653 | -104.1409
26.6567 | -104.1444
26.6605 | -104.1493
26.6642 | -104.154
26.6679 | -104.1562
26.6716 | -104.1587
26.6754 | -104.1606
26.6791 | -104.1623
26.6828 | -104.165
26.6865 | -104.1684
26.6903 | -104.1729
26.694 | -104.176
26.6977 | -104.1773
26.7015 | -104.1787
26.7052 | -104.182
26.7089 | -104.185
26.7126 | -104.1855
26.7164 | -104.1887
26.7201 | -104.1926
26.7238 | -104.1957
26.7275 | -104.1977
26.7313 | -104.1993
26.735 | -104.201
26.7387 | -104.2039
26.7425 | -104.2074
26.7462 | -104.2121
26.7499 | -104.2169
26.7536 | -104.2212
26.7574 | -104.2253
26.7611 | -104.2298
26.7648 | -104.2359
26.7685 | -104.2405
26.7723 | -104.2458
26.776 | -104.2522
26.7797 | -104.2577
26.7835 | -104.2642
26.7872 | -104.2697
26.7909 | -104.2759
26.7946 | -104.2827
26.7984 | -104.2903
26.8021 | -104.2983
26.8058 | -104.3051
26.8095 | -104.3118
26.8133 | -104.32
26.817 | -104.3282
26.8207 | -104.3365
26.8244 | -104.3449
26.8282 | -104.3535
26.8319 | -104.3612
26.8356 | -104.3698
26.8394 | -104.3794
26.8431 | -104.389
26.8468 | -104.3989
26.8505 | -104.4087
26.8543 | -104.4171
26.858 | -104.4234
26.8617 | -104.4302
26.8654 | -104.4364
26.8692 | -104.4421
26.8729 | -104.4487
26.8766 | -104.4554
26.8804 | -104.4613
26.8841 | -104.4648
26.8878 | -104.4706
26.8915 | -104.4745
26.8953 | -104.4789
26.899 | -104.4841
26.9027 | -104.489
26.9064 | -104.4927
26.9102 | -104.4977
26.9139 | -104.5019
26.9176 | -104.5049
26.9214 | -104.5048
26.9251 | -104.5056
26.9288 | -104.5076
26.9325 | -104.5092
26.9363 | -104.5105
26.94 | -104.5107
26.9437 | -104.5107
26.9474 | -104.5114
26.9512 | -104.5113
26.9549 | -104.5095
26.9586 | -104.5066
26.9624 | -104.5032
26.9661 | -104.4993
26.9698 | -104.4945
26.9735 | -104.4893
26.9773 | -104.4863
26.981 | -104.486
26.9847 | -104.4836
26.9884 | -104.4784
26.9922 | -104.4749
26.9959 | -104.4704
26.9996 | -104.4649
27.0034 | -104.4588
27.0071 | -104.453
27.0108 | -104.4478
27.0145 | -104.4427
27.0183 | -104.4376
27.022 | -104.4316
27.0257 | -104.4248
27.0294 | -104.4191
27.0332 | -104.4142
27.0369 | -104.4074
27.0406 | -104.4011
27.0444 | -104.3953
27.0481 | -104.3915
27.0518 | -104.3849
27.0555 | -104.3766
27.0593 | -104.3685
27.063 | -104.3618
27.0667 | -104.3563
27.0704 | -104.3508
27.0742 | -104.3459
27.0779 | -104.3406
27.0816 | -104.3355
27.0854 | -104.3297
27.0891 | -104.3255
27.0928 | -104.3216
27.0965 | -104.315
27.1003 | -104.3101
27.104 | -104.3064
27.1077 | -104.3034
27.1114 | -104.301
27.1152 | -104.3
27.1189 | -104.2982
27.1226 | -104.2959
27.1264 | -104.2913
27.1301 | -104.2887
27.1338 | -104.2861
27.1375 | -104.2839
27.1413 | -104.2828
27.145 | -104.2818
27.1487 | -104.2806
27.1524 | -104.2791
27.1562 | -104.2787
27.1599 | -104.2799
27.1636 | -104.2793
27.1673 | -104.2795
27.1711 | -104.2798
27.1748 | -104.2791
27.1785 | -104.2785
27.1823 | -104.2793
27.186 | -104.2797
27.1897 | -104.2805
27.1934 | -104.2814
27.1972 | -104.2824
27.2009 | -104.2837
27.2046 | -104.2852
27.2083 | -104.285
27.2121 | -104.2864
27.2158 | -104.2893
27.2195 | -104.2909
27.2233 | -104.2912
27.227 | -104.2915
27.2307 | -104.2926
27.2344 | -104.2968
27.2382 | -104.3011
27.2419 | -104.3019
27.2456 | -104.3036
27.2493 | -104.3073
27.2531 | -104.3097
27.2568 | -104.3126
27.2605 | -104.3159
27.2643 | -104.3183
27.268 | -104.3207
27.2717 | -104.3235
27.2754 | -104.3266
27.2792 | -104.3273
27.2829 | -104.3267
27.2866 | -104.3279
27.2903 | -104.3301
27.2941 | -104.3299
27.2978 | -104.3299
27.3015 | -104.3302
27.3053 | -104.3311
27.309 | -104.3313
27.3127 | -104.3302
27.3164 | -104.3279
27.3202 | -104.3275
27.3239 | -104.3275
27.3276 | -104.326
27.3313 | -104.3229
27.3351 | -104.319
27.3388 | -104.3173
27.3425 | -104.3168
27.3463 | -104.3152
27.35 | -104.3115
27.3537 | -104.307
27.3574 | -104.3032
27.3612 | -104.2972
27.3649 | -104.2902
27.3686 | -104.2831
27.3723 | -104.2763
27.3761 | -104.2686
27.3798 | -104.2602
27.3835 | -104.2521
27.3873 | -104.2445
27.391 | -104.238
27.3947 | -104.2293
27.3984 | -104.2186
27.4022 | -104.2081
27.4059 | -104.199
27.4096 | -104.1911
27.4133 | -104.1813
27.4171 | -104.1713
27.4208 | -104.16
27.4245 | -104.1486
27.4283 | -104.1371
27.432 | -104.1254
27.4357 | -104.1144
27.4394 | -104.1058
27.4432 | -104.0966
27.4469 | -104.0877
27.4506 | -104.0776
27.4543 | -104.0665
27.4581 | -104.056
27.4618 | -104.0467
27.4655 | -104.0372
27.4693 | -104.0275
27.473 | -104.0185
27.4767 | -104.0103
27.4804 | -104.0016
27.4842 | -103.9931
27.4879 | -103.9856
27.4916 | -103.9778
27.4953 | -103.9689
27.4991 | -103.9608
27.5028 | -103.9537
27.5065 | -103.9476
27.5102 | -103.9421
27.514 | -103.9392
27.5177 | -103.9355
27.5214 | -103.9315
27.5252 | -103.928
27.5289 | -103.925
27.5326 | -103.9222
27.5363 | -103.921
27.5401 | -103.9209
27.5438 | -103.9202
27.5475 | -103.9196
27.5512 | -103.9199
27.555 | -103.9203
27.5587 | -103.9212
27.5624 | -103.9227
27.5662 | -103.9246
27.5699 | -103.9264
27.5736 | -103.928
27.5773 | -103.9294
27.5811 | -103.9332
27.5848 | -103.9373
27.5885 | -103.9407
27.5922 | -103.9436
27.596 | -103.9493
27.5997 | -103.9542
27.6034 | -103.9586
27.6072 | -103.9632
27.6109 | -103.9682
27.6146 | -103.9731
27.6183 | -103.9762
27.6221 | -103.9795
27.6258 | -103.9833
27.6295 | -103.9885
27.6332 | -103.9939
27.637 | -103.999
27.6407 | -104.0052
27.6444 | -104.0115
27.6482 | -104.0167
27.6519 | -104.0214
27.6556 | -104.0263
27.6593 | -104.0308
27.6631 | -104.0363
27.6668 | -104.0416
27.6705 | -104.047
27.6742 | -104.0508
27.678 | -104.0544
27.6817 | -104.0576
27.6854 | -104.0598
27.6892 | -104.0627
27.6929 | -104.0664
27.6966 | -104.0672
27.7003 | -104.0688
27.7041 | -104.0708
27.7078 | -104.0706
27.7115 | -104.0701
27.7152 | -104.0716
27.719 | -104.0735
27.7227 | -104.0755
27.7264 | -104.0759
27.7302 | -104.0751
27.7339 | -104.0743
27.7376 | -104.0741
27.7413 | -104.0736
27.7451 | -104.0724
27.7488 | -104.0717
27.7525 | -104.0716
27.7562 | -104.0721
27.76 | -104.0725
27.7637 | -104.0726
27.7674 | -104.0721
27.7712 | -104.073
27.7749 | -104.0729
27.7786 | -104.0712
27.7823 | -104.0705
27.7861 | -104.0708
27.7898 | -104.0719
27.7935 | -104.0729
27.7972 | -104.073
27.801 | -104.0751
27.8047 | -104.077
27.8084 | -104.0772
27.8122 | -104.0775
27.8159 | -104.0816
27.8196 | -104.0848
27.8233 | -104.0885
27.8271 | -104.0926
27.8308 | -104.0957
27.8345 | -104.1011
27.8382 | -104.1057
27.842 | -104.1077
27.8457 | -104.1121
27.8494 | -104.1161
27.8531 | -104.1204
27.8569 | -104.1266
27.8606 | -104.1323
27.8643 | -104.1373
27.8681 | -104.1409
27.8718 | -104.1443
27.8755 | -104.1474
27.8792 | -104.149
27.883 | -104.1512
27.8867 | -104.1571
27.8904 | -104.162
27.8941 | -104.1654
27.8979 | -104.1689
27.9016 | -104.1727
27.9053 | -104.1763
27.9091 | -104.1809
27.9128 | -104.1844
27.9165 | -104.1862
27.9202 | -104.1889
27.924 | -104.1921
27.9277 | -104.1938
27.9314 | -104.1922
27.9351 | -104.1915
27.9389 | -104.1913
27.9426 | -104.1916
27.9463 | -104.192
27.9501 | -104.1927
27.9538 | -104.1939
27.9575 | -104.1899
27.9612 | -104.1864
27.965 | -104.1828
27.9687 | -104.1799
27.9724 | -104.1778
27.9761 | -104.1727
27.9799 | -104.1684
27.9836 | -104.1654
27.9873 | -104.1626
27.9911 | -104.1569
27.9948 | -104.1506
27.9985 | -104.1446
28.0022 | -104.1389