## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 236 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 238
Hood | 02-92B/P | 177
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1009.3 | hPa |  | PRC PRO 01
Humidity | 51.6 | %/Hr
Internal temperature of calibration bench | 23.24 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1016 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.3 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.04 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.03 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3060 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 24/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 27/05/2018
Engineer | Anthony HUE | Approval |  | 29/05/2018
CEO | Robert GLEMOT | Approval |  | 29/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4346 | 57.6056
0.012209 | -35.9851 | 38.3758
0.018313 | -34.9131 | 28.2503
0.024417 | -34.4516 | 22.0348
0.030521 | -34.3453 | 18.2881
0.036626 | -34.2016 | 15.2232
0.04273 | -34.1048 | 12.83
0.048834 | -34.0578 | 11.3249
0.054938 | -33.9891 | 10.1954
0.061043 | -33.9248 | 9.1033
0.067147 | -33.9146 | 8.2693
0.073251 | -33.9166 | 7.5017
0.079355 | -33.9373 | 6.8395
0.08546 | -33.9532 | 6.2448
0.091564 | -33.9483 | 5.8078
0.097668 | -33.9284 | 5.5714
0.10377 | -33.9043 | 5.3283
0.10988 | -33.8856 | 5.0525
0.11598 | -33.8758 | 4.7572
0.12209 | -33.8781 | 4.4696
0.12819 | -33.8894 | 4.1809
0.13429 | -33.9016 | 3.8874
0.1404 | -33.9048 | 3.6041
0.1465 | -33.9066 | 3.342
0.15261 | -33.9032 | 3.1184
0.15871 | -33.8958 | 2.9553
0.16482 | -33.8869 | 2.8542
0.17092 | -33.8846 | 2.8206
0.17702 | -33.8823 | 2.7518
0.18313 | -33.8815 | 2.6768
0.18923 | -33.881 | 2.5685
0.19534 | -33.8795 | 2.441
0.20144 | -33.8777 | 2.2964
0.20754 | -33.8741 | 2.1735
0.21365 | -33.8727 | 2.0576
0.21975 | -33.8719 | 1.9593
0.22586 | -33.8729 | 1.8722
0.23196 | -33.8754 | 1.8012
0.23807 | -33.8798 | 1.7402
0.24417 | -33.8843 | 1.6822
0.25027 | -33.8886 | 1.6202
0.25638 | -33.8912 | 1.5582
0.26248 | -33.8932 | 1.4905
0.26859 | -33.8942 | 1.4238
0.27469 | -33.8947 | 1.3509
0.2808 | -33.8958 | 1.2758
0.2869 | -33.8968 | 1.2037
0.293 | -33.8978 | 1.1359
0.29911 | -33.8989 | 1.0758
0.30521 | -33.8999 | 1.0167
0.31132 | -33.9007 | 0.97962
0.31742 | -33.902 | 0.93798
0.32353 | -33.9023 | 0.90463
0.32963 | -33.9032 | 0.87762
0.33573 | -33.9042 | 0.85975
0.34184 | -33.9057 | 0.84036
0.34794 | -33.908 | 0.81778
0.35405 | -33.9104 | 0.79507
0.36015 | -33.9131 | 0.77074
0.36626 | -33.9153 | 0.74389
0.37236 | -33.9177 | 0.70691
0.37846 | -33.92 | 0.66149
0.38457 | -33.9209 | 0.60196
0.39067 | -33.922 | 0.554
0.39678 | -33.9227 | 0.50033
0.40288 | -33.9231 | 0.44414
0.40899 | -33.9229 | 0.3904
0.41509 | -33.922 | 0.33783
0.42119 | -33.9218 | 0.29357
0.4273 | -33.9216 | 0.24601
0.4334 | -33.9214 | 0.1985
0.43951 | -33.9215 | 0.15401
0.44561 | -33.9223 | 0.11343
0.45172 | -33.9239 | 0.081343
0.45782 | -33.9246 | 0.043503
0.46392 | -33.9258 | 0.0064121
0.47003 | -33.9269 | -0.026441
0.47613 | -33.9277 | -0.061304
0.48224 | -33.9282 | -0.097027
0.48834 | -33.9272 | -0.13287
0.49445 | -33.9269 | -0.16915
0.50055 | -33.9262 | -0.20202
0.50665 | -33.9253 | -0.23731
0.51276 | -33.9242 | -0.27244
0.51886 | -33.923 | -0.30736
0.52497 | -33.9217 | -0.33165
0.53107 | -33.9208 | -0.36174
0.53717 | -33.92 | -0.39045
0.54328 | -33.919 | -0.41335
0.54938 | -33.9188 | -0.43576
0.55549 | -33.9187 | -0.45648
0.56159 | -33.9199 | -0.46706
0.5677 | -33.9202 | -0.48086
0.5738 | -33.9213 | -0.49657
0.5799 | -33.9221 | -0.5108
0.58601 | -33.9233 | -0.51971
0.59211 | -33.925 | -0.52451
0.59822 | -33.9265 | -0.5328
0.60432 | -33.9281 | -0.54401
0.61043 | -33.9294 | -0.55057
0.61653 | -33.9307 | -0.55977
0.62263 | -33.9323 | -0.56798
0.62874 | -33.9329 | -0.57756
0.63484 | -33.9338 | -0.58563
0.64095 | -33.9343 | -0.5948
0.64705 | -33.9346 | -0.60349
0.65316 | -33.9347 | -0.61513
0.65926 | -33.9348 | -0.63052
0.66536 | -33.9335 | -0.64786
0.67147 | -33.933 | -0.66049
0.67757 | -33.9325 | -0.67431
0.68368 | -33.9316 | -0.69328
0.68978 | -33.9307 | -0.71038
0.69589 | -33.9296 | -0.72949
0.70199 | -33.9286 | -0.75286
0.70809 | -33.9275 | -0.77225
0.7142 | -33.9267 | -0.79309
0.7203 | -33.9255 | -0.81255
0.72641 | -33.9247 | -0.83563
0.73251 | -33.924 | -0.85601
0.73862 | -33.9242 | -0.87839
0.74472 | -33.9238 | -0.89881
0.75082 | -33.9235 | -0.9215
0.75693 | -33.9238 | -0.94389
0.76303 | -33.9244 | -0.96646
0.76914 | -33.9264 | -0.98724
0.77524 | -33.9274 | -1.0072
0.78135 | -33.9287 | -1.0276
0.78745 | -33.9302 | -1.0478
0.79355 | -33.9315 | -1.0679
0.79966 | -33.9334 | -1.0876
0.80576 | -33.9359 | -1.1074
0.81187 | -33.9376 | -1.1294
0.81797 | -33.9397 | -1.151
0.82408 | -33.9415 | -1.1719
0.83018 | -33.9434 | -1.1927
0.83628 | -33.9454 | -1.2112
0.84239 | -33.9467 | -1.2321
0.84849 | -33.9482 | -1.2542
0.8546 | -33.9498 | -1.2742
0.8607 | -33.9511 | -1.2952
0.86681 | -33.9523 | -1.3143
0.87291 | -33.9533 | -1.3333
0.87901 | -33.9538 | -1.3515
0.88512 | -33.9546 | -1.3698
0.89122 | -33.9554 | -1.3882
0.89733 | -33.9558 | -1.4037
0.90343 | -33.9562 | -1.4214
0.90953 | -33.9564 | -1.4357
0.91564 | -33.9564 | -1.4515
0.92174 | -33.9566 | -1.4664
0.92785 | -33.9567 | -1.4808
0.93395 | -33.9567 | -1.4945
0.94006 | -33.9569 | -1.5071
0.94616 | -33.9566 | -1.5186
0.95226 | -33.9567 | -1.531
0.95837 | -33.9568 | -1.5447
0.96447 | -33.9569 | -1.5573
0.97058 | -33.9567 | -1.5705
0.97668 | -33.9565 | -1.5837
0.98279 | -33.9562 | -1.5962
0.98889 | -33.9559 | -1.6064
0.99499 | -33.9559 | -1.6197
1.0011 | -33.9557 | -1.6299
1.0072 | -33.9553 | -1.6412
1.0133 | -33.9549 | -1.6515
1.0194 | -33.9544 | -1.6643
1.0255 | -33.954 | -1.6779
1.0316 | -33.9534 | -1.6894
1.0377 | -33.9529 | -1.7029
1.0438 | -33.9522 | -1.7154
1.0499 | -33.9513 | -1.7306
1.056 | -33.9505 | -1.7448
1.0621 | -33.9498 | -1.759
1.0682 | -33.9492 | -1.7711
1.0743 | -33.9485 | -1.7866
1.0805 | -33.9479 | -1.8009
1.0866 | -33.9468 | -1.8159
1.0927 | -33.9459 | -1.8303
1.0988 | -33.9451 | -1.8451
1.1049 | -33.9442 | -1.8592
1.111 | -33.9433 | -1.8746
1.1171 | -33.9425 | -1.889
1.1232 | -33.9419 | -1.9042
1.1293 | -33.9414 | -1.9181
1.1354 | -33.941 | -1.9331
1.1415 | -33.9406 | -1.9486
1.1476 | -33.9399 | -1.9642
1.1537 | -33.9393 | -1.9788
1.1598 | -33.9393 | -1.9935
1.1659 | -33.939 | -2.0066
1.172 | -33.9387 | -2.0205
1.1781 | -33.9384 | -2.0344
1.1842 | -33.9382 | -2.0485
1.1903 | -33.9386 | -2.0591
1.1964 | -33.9384 | -2.0724
1.2025 | -33.9385 | -2.0863
1.2086 | -33.9388 | -2.0985
1.2147 | -33.9388 | -2.1108
1.2209 | -33.9392 | -2.1217
1.227 | -33.9399 | -2.1322
1.2331 | -33.9404 | -2.1439
1.2392 | -33.9407 | -2.1555
1.2453 | -33.9412 | -2.1663
1.2514 | -33.9416 | -2.1758
1.2575 | -33.9422 | -2.1865
1.2636 | -33.9431 | -2.1949
1.2697 | -33.9438 | -2.2037
1.2758 | -33.9445 | -2.2146
1.2819 | -33.9453 | -2.2241
1.288 | -33.946 | -2.2346
1.2941 | -33.9469 | -2.2425
1.3002 | -33.9476 | -2.25
1.3063 | -33.9482 | -2.2586
1.3124 | -33.9489 | -2.2671
1.3185 | -33.9497 | -2.2757
1.3246 | -33.9504 | -2.2844
1.3307 | -33.9513 | -2.2925
1.3368 | -33.9522 | -2.3006
1.3429 | -33.9528 | -2.3091
1.349 | -33.9534 | -2.318
1.3551 | -33.9541 | -2.3269
1.3613 | -33.9547 | -2.3347
1.3674 | -33.9554 | -2.3443
1.3735 | -33.9558 | -2.3537
1.3796 | -33.9563 | -2.3622
1.3857 | -33.9567 | -2.371
1.3918 | -33.9572 | -2.3808
1.3979 | -33.9578 | -2.3895
1.404 | -33.9583 | -2.4006
1.4101 | -33.9588 | -2.4106
1.4162 | -33.9592 | -2.4208
1.4223 | -33.9597 | -2.4319
1.4284 | -33.9602 | -2.4415
1.4345 | -33.9604 | -2.4528
1.4406 | -33.9607 | -2.4662
1.4467 | -33.9612 | -2.4776
1.4528 | -33.9617 | -2.4886
1.4589 | -33.9622 | -2.5003
1.465 | -33.9624 | -2.5108
1.4711 | -33.9629 | -2.5219
1.4772 | -33.9635 | -2.5352
1.4833 | -33.9639 | -2.5481
1.4894 | -33.9645 | -2.5601
1.4955 | -33.965 | -2.5718
1.5016 | -33.9656 | -2.5842
1.5078 | -33.966 | -2.5981
1.5139 | -33.9664 | -2.611
1.52 | -33.9669 | -2.6234
1.5261 | -33.9672 | -2.6359
1.5322 | -33.9675 | -2.6499
1.5383 | -33.9679 | -2.6623
1.5444 | -33.9685 | -2.6748
1.5505 | -33.969 | -2.6882
1.5566 | -33.9695 | -2.7011
1.5627 | -33.97 | -2.7148
1.5688 | -33.9706 | -2.7284
1.5749 | -33.971 | -2.7409
1.581 | -33.9713 | -2.7542
1.5871 | -33.9717 | -2.7679
1.5932 | -33.9721 | -2.7807
1.5993 | -33.9725 | -2.7949
1.6054 | -33.9728 | -2.8081
1.6115 | -33.9733 | -2.8222
1.6176 | -33.9736 | -2.8347
1.6237 | -33.9738 | -2.8483
1.6298 | -33.974 | -2.8607
1.6359 | -33.9744 | -2.8734
1.642 | -33.9748 | -2.8851
1.6482 | -33.975 | -2.898
1.6543 | -33.9754 | -2.9117
1.6604 | -33.9757 | -2.9236
1.6665 | -33.9758 | -2.9359
1.6726 | -33.976 | -2.9483
1.6787 | -33.9763 | -2.9607
1.6848 | -33.9762 | -2.9731
1.6909 | -33.9764 | -2.9857
1.697 | -33.9767 | -2.9977
1.7031 | -33.9767 | -3.0102
1.7092 | -33.9766 | -3.0221
1.7153 | -33.9768 | -3.0339
1.7214 | -33.9765 | -3.046
1.7275 | -33.9764 | -3.0584
1.7336 | -33.9763 | -3.0713
1.7397 | -33.9764 | -3.0834
1.7458 | -33.9764 | -3.0954
1.7519 | -33.9764 | -3.1074
1.758 | -33.9764 | -3.1191
1.7641 | -33.9765 | -3.1305
1.7702 | -33.9764 | -3.1433
1.7763 | -33.9763 | -3.1545
1.7824 | -33.9763 | -3.1668
1.7885 | -33.9762 | -3.1791
1.7947 | -33.9763 | -3.1909
1.8008 | -33.9763 | -3.2019
1.8069 | -33.9762 | -3.2132
1.813 | -33.9763 | -3.2253
1.8191 | -33.9762 | -3.2379
1.8252 | -33.9763 | -3.2505
1.8313 | -33.9764 | -3.2626
1.8374 | -33.9765 | -3.2743
1.8435 | -33.9765 | -3.286
1.8496 | -33.9766 | -3.299
1.8557 | -33.9767 | -3.3117
1.8618 | -33.977 | -3.3238
1.8679 | -33.9771 | -3.3357
1.874 | -33.9771 | -3.3475
1.8801 | -33.9773 | -3.3598
1.8862 | -33.9775 | -3.3716
1.8923 | -33.9776 | -3.3837
1.8984 | -33.9777 | -3.3957
1.9045 | -33.9778 | -3.4071
1.9106 | -33.9778 | -3.4192
1.9167 | -33.9778 | -3.4313
1.9228 | -33.9779 | -3.443
1.9289 | -33.978 | -3.4548
1.9351 | -33.9782 | -3.4654
1.9412 | -33.9784 | -3.4765
1.9473 | -33.9787 | -3.4885
1.9534 | -33.9789 | -3.4998
1.9595 | -33.979 | -3.5115
1.9656 | -33.9793 | -3.5238
1.9717 | -33.9795 | -3.5354
1.9778 | -33.9799 | -3.5466
1.9839 | -33.9801 | -3.558
1.99 | -33.9803 | -3.5689
1.9961 | -33.9804 | -3.581
2.0022 | -33.9806 | -3.5924
2.0083 | -33.9807 | -3.6034
2.0144 | -33.9809 | -3.6147
2.0205 | -33.9812 | -3.6264
2.0266 | -33.9815 | -3.6377
2.0327 | -33.9818 | -3.6492
2.0388 | -33.9822 | -3.6607
2.0449 | -33.9824 | -3.672
2.051 | -33.9827 | -3.6832
2.0571 | -33.983 | -3.6945
2.0632 | -33.9832 | -3.7058
2.0693 | -33.9834 | -3.718
2.0754 | -33.9836 | -3.7292
2.0816 | -33.984 | -3.7413
2.0877 | -33.9842 | -3.7527
2.0938 | -33.9845 | -3.7648
2.0999 | -33.9847 | -3.7765
2.106 | -33.9849 | -3.788
2.1121 | -33.9851 | -3.7999
2.1182 | -33.9854 | -3.8105
2.1243 | -33.9856 | -3.821
2.1304 | -33.9858 | -3.8321
2.1365 | -33.9861 | -3.8438
2.1426 | -33.9864 | -3.8556
2.1487 | -33.9867 | -3.8667
2.1548 | -33.9871 | -3.8785
2.1609 | -33.9876 | -3.8898
2.167 | -33.988 | -3.9009
2.1731 | -33.9883 | -3.9129
2.1792 | -33.9885 | -3.9237
2.1853 | -33.9889 | -3.9356
2.1914 | -33.9892 | -3.9469
2.1975 | -33.9896 | -3.9583
2.2036 | -33.9898 | -3.9704
2.2097 | -33.9901 | -3.9817
2.2158 | -33.9905 | -3.9922
2.222 | -33.9907 | -4.0024
2.2281 | -33.991 | -4.0137
2.2342 | -33.9913 | -4.0249
2.2403 | -33.9915 | -4.0354
2.2464 | -33.9918 | -4.0459
2.2525 | -33.992 | -4.0572
2.2586 | -33.9924 | -4.0676
2.2647 | -33.9927 | -4.0777
2.2708 | -33.993 | -4.0888
2.2769 | -33.9934 | -4.0993
2.283 | -33.9935 | -4.1103
2.2891 | -33.9938 | -4.121
2.2952 | -33.9939 | -4.1313
2.3013 | -33.9942 | -4.142
2.3074 | -33.9945 | -4.1523
2.3135 | -33.9947 | -4.1635
2.3196 | -33.9949 | -4.1735
2.3257 | -33.9951 | -4.1839
2.3318 | -33.9952 | -4.1949
2.3379 | -33.9954 | -4.2054
2.344 | -33.9955 | -4.2167
2.3501 | -33.9957 | -4.2272
2.3562 | -33.9959 | -4.2374
2.3623 | -33.996 | -4.2475
2.3685 | -33.9961 | -4.2578
2.3746 | -33.9962 | -4.2689
2.3807 | -33.9965 | -4.2791
2.3868 | -33.9966 | -4.2896
2.3929 | -33.9968 | -4.3001
2.399 | -33.9969 | -4.3104
2.4051 | -33.9972 | -4.3207
2.4112 | -33.9973 | -4.3311
2.4173 | -33.9975 | -4.3415
2.4234 | -33.9975 | -4.3521
2.4295 | -33.9977 | -4.3628
2.4356 | -33.9979 | -4.3734
2.4417 | -33.998 | -4.3844
2.4478 | -33.9981 | -4.3951
2.4539 | -33.9983 | -4.4051
2.46 | -33.9983 | -4.4164
2.4661 | -33.9985 | -4.4273
2.4722 | -33.9986 | -4.4372
2.4783 | -33.9987 | -4.4474
2.4844 | -33.9988 | -4.4583
2.4905 | -33.9988 | -4.4691
2.4966 | -33.999 | -4.4797
2.5027 | -33.9991 | -4.4899
2.5089 | -33.9992 | -4.5006
2.515 | -33.9992 | -4.5113
2.5211 | -33.9993 | -4.5218
2.5272 | -33.9993 | -4.5325
2.5333 | -33.9994 | -4.5433
2.5394 | -33.9996 | -4.5539
2.5455 | -33.9997 | -4.5651
2.5516 | -33.9998 | -4.5752
2.5577 | -34 | -4.5865
2.5638 | -34.0001 | -4.5981
2.5699 | -34.0002 | -4.6094
2.576 | -34.0002 | -4.6208
2.5821 | -34.0003 | -4.6319
2.5882 | -34.0004 | -4.6429
2.5943 | -34.0005 | -4.6535
2.6004 | -34.0007 | -4.6643
2.6065 | -34.0008 | -4.6745
2.6126 | -34.0009 | -4.6847
2.6187 | -34.001 | -4.695
2.6248 | -34.0011 | -4.7055
2.6309 | -34.0013 | -4.7168
2.637 | -34.0015 | -4.7283
2.6431 | -34.0016 | -4.7386
2.6492 | -34.0017 | -4.7498
2.6554 | -34.0019 | -4.7605
2.6615 | -34.002 | -4.7713
2.6676 | -34.0022 | -4.7825
2.6737 | -34.0025 | -4.7935
2.6798 | -34.0028 | -4.8045
2.6859 | -34.003 | -4.8156
2.692 | -34.0031 | -4.8263
2.6981 | -34.0033 | -4.8372
2.7042 | -34.0035 | -4.8482
2.7103 | -34.0036 | -4.8586
2.7164 | -34.004 | -4.8694
2.7225 | -34.0042 | -4.8806
2.7286 | -34.0044 | -4.8915
2.7347 | -34.0046 | -4.9018
2.7408 | -34.0049 | -4.9116
2.7469 | -34.0051 | -4.9228
2.753 | -34.0054 | -4.9332
2.7591 | -34.0057 | -4.9443
2.7652 | -34.0059 | -4.9557
2.7713 | -34.006 | -4.9662
2.7774 | -34.0063 | -4.9769
2.7835 | -34.0065 | -4.9877
2.7896 | -34.0067 | -4.998
2.7958 | -34.0069 | -5.0086
2.8019 | -34.0072 | -5.0188
2.808 | -34.0075 | -5.0293
2.8141 | -34.0078 | -5.0403
2.8202 | -34.0081 | -5.051
2.8263 | -34.0084 | -5.0617
2.8324 | -34.0088 | -5.073
2.8385 | -34.0092 | -5.0836
2.8446 | -34.0095 | -5.0942
2.8507 | -34.0099 | -5.1049
2.8568 | -34.0103 | -5.1151
2.8629 | -34.0105 | -5.1255
2.869 | -34.0109 | -5.1364
2.8751 | -34.0112 | -5.147
2.8812 | -34.0116 | -5.1582
2.8873 | -34.0119 | -5.1689
2.8934 | -34.0122 | -5.1797
2.8995 | -34.0126 | -5.1907
2.9056 | -34.0128 | -5.2017
2.9117 | -34.0132 | -5.2126
2.9178 | -34.0136 | -5.2231
2.9239 | -34.014 | -5.2346
2.93 | -34.0144 | -5.2459
2.9361 | -34.0147 | -5.2566
2.9423 | -34.015 | -5.2671
2.9484 | -34.0155 | -5.278
2.9545 | -34.0159 | -5.2889
2.9606 | -34.0163 | -5.3001
2.9667 | -34.0166 | -5.3112
2.9728 | -34.017 | -5.3223
2.9789 | -34.0174 | -5.3333
2.985 | -34.0179 | -5.3445
2.9911 | -34.0183 | -5.3557
2.9972 | -34.0188 | -5.3666
3.0033 | -34.0192 | -5.3781
3.0094 | -34.0196 | -5.389
3.0155 | -34.02 | -5.3998
3.0216 | -34.0204 | -5.411
3.0277 | -34.0207 | -5.4225
3.0338 | -34.0211 | -5.4339
3.0399 | -34.0216 | -5.4451
3.046 | -34.0219 | -5.4566
3.0521 | -34.0222 | -5.4676
3.0582 | -34.0225 | -5.4796
3.0643 | -34.0229 | -5.491
3.0704 | -34.0232 | -5.5028
3.0765 | -34.0235 | -5.5138
3.0827 | -34.0238 | -5.5249
3.0888 | -34.0242 | -5.5361
3.0949 | -34.0246 | -5.5474
3.101 | -34.0249 | -5.5588
3.1071 | -34.0252 | -5.5705
3.1132 | -34.0256 | -5.5822
3.1193 | -34.026 | -5.5936
3.1254 | -34.0264 | -5.6049
3.1315 | -34.0268 | -5.6168
3.1376 | -34.027 | -5.6281
3.1437 | -34.0274 | -5.6396
3.1498 | -34.0278 | -5.6517
3.1559 | -34.0281 | -5.663
3.162 | -34.0284 | -5.6752
3.1681 | -34.0286 | -5.6871
3.1742 | -34.029 | -5.6986
3.1803 | -34.0293 | -5.7102
3.1864 | -34.0297 | -5.7214
3.1925 | -34.03 | -5.7329
3.1986 | -34.0302 | -5.7449
3.2047 | -34.0305 | -5.757
3.2108 | -34.0308 | -5.768
3.2169 | -34.0311 | -5.7797
3.223 | -34.0313 | -5.7914
3.2292 | -34.0316 | -5.8032
3.2353 | -34.0318 | -5.8139
3.2414 | -34.032 | -5.8253
3.2475 | -34.0322 | -5.8367
3.2536 | -34.0325 | -5.8481
3.2597 | -34.0328 | -5.8589
3.2658 | -34.033 | -5.8698
3.2719 | -34.0332 | -5.8807
3.278 | -34.0334 | -5.8923
3.2841 | -34.0337 | -5.9033
3.2902 | -34.0339 | -5.9147
3.2963 | -34.0342 | -5.9262
3.3024 | -34.0343 | -5.9372
3.3085 | -34.0346 | -5.9487
3.3146 | -34.0347 | -5.9597
3.3207 | -34.035 | -5.9712
3.3268 | -34.0352 | -5.9826
3.3329 | -34.0354 | -5.9933
3.339 | -34.0355 | -6.004
3.3451 | -34.0357 | -6.0151
3.3512 | -34.0359 | -6.0257
3.3573 | -34.0361 | -6.0369
3.3634 | -34.0363 | -6.048
3.3696 | -34.0365 | -6.0593
3.3757 | -34.0366 | -6.07
3.3818 | -34.0368 | -6.0808
3.3879 | -34.037 | -6.0918
3.394 | -34.0371 | -6.1031
3.4001 | -34.0373 | -6.1138
3.4062 | -34.0374 | -6.1247
3.4123 | -34.0375 | -6.1351
3.4184 | -34.0377 | -6.1463
3.4245 | -34.0379 | -6.157
3.4306 | -34.038 | -6.1676
3.4367 | -34.0382 | -6.178
3.4428 | -34.0383 | -6.1882
3.4489 | -34.0384 | -6.1991
3.455 | -34.0386 | -6.2099
3.4611 | -34.0388 | -6.221
3.4672 | -34.0389 | -6.232
3.4733 | -34.039 | -6.243
3.4794 | -34.0392 | -6.2537
3.4855 | -34.0393 | -6.264
3.4916 | -34.0395 | -6.2742
3.4977 | -34.0396 | -6.2854
3.5038 | -34.0398 | -6.2962
3.5099 | -34.0399 | -6.3063
3.5161 | -34.04 | -6.3165
3.5222 | -34.0401 | -6.3267
3.5283 | -34.0402 | -6.3374
3.5344 | -34.0403 | -6.3481
3.5405 | -34.0404 | -6.3585
3.5466 | -34.0406 | -6.3694
3.5527 | -34.0407 | -6.38
3.5588 | -34.0408 | -6.3905
3.5649 | -34.0408 | -6.401
3.571 | -34.0409 | -6.4115
3.5771 | -34.041 | -6.4225
3.5832 | -34.0412 | -6.4322
3.5893 | -34.0414 | -6.4427
3.5954 | -34.0415 | -6.4535
3.6015 | -34.0417 | -6.4642
3.6076 | -34.0419 | -6.4746
3.6137 | -34.042 | -6.4845
3.6198 | -34.0423 | -6.4946
3.6259 | -34.0424 | -6.5054
3.632 | -34.0425 | -6.5159
3.6381 | -34.0427 | -6.5266
3.6442 | -34.0428 | -6.537
3.6503 | -34.043 | -6.547
3.6565 | -34.0432 | -6.5576
3.6626 | -34.0433 | -6.5681
3.6687 | -34.0435 | -6.5787
3.6748 | -34.0438 | -6.589
3.6809 | -34.044 | -6.5995
3.687 | -34.0441 | -6.6103
3.6931 | -34.0444 | -6.6204
3.6992 | -34.0446 | -6.6309
3.7053 | -34.0448 | -6.6412
3.7114 | -34.0449 | -6.6518
3.7175 | -34.0451 | -6.6623
3.7236 | -34.0454 | -6.6722
3.7297 | -34.0456 | -6.6829
3.7358 | -34.0458 | -6.6937
3.7419 | -34.0459 | -6.7044
3.748 | -34.0462 | -6.7148
3.7541 | -34.0463 | -6.7253
3.7602 | -34.0466 | -6.7356
3.7663 | -34.0468 | -6.7459
3.7724 | -34.0471 | -6.7565
3.7785 | -34.0473 | -6.7667
3.7846 | -34.0475 | -6.777
3.7907 | -34.0478 | -6.7878
3.7969 | -34.0481 | -6.7984
3.803 | -34.0483 | -6.8084
3.8091 | -34.0485 | -6.8188
3.8152 | -34.0488 | -6.8292
3.8213 | -34.049 | -6.8399
3.8274 | -34.0493 | -6.8501
3.8335 | -34.0496 | -6.8607
3.8396 | -34.0499 | -6.8712
3.8457 | -34.0502 | -6.8818
3.8518 | -34.0506 | -6.8922
3.8579 | -34.0509 | -6.9026
3.864 | -34.0512 | -6.9132
3.8701 | -34.0515 | -6.9241
3.8762 | -34.0517 | -6.9344
3.8823 | -34.052 | -6.9451
3.8884 | -34.0523 | -6.9556
3.8945 | -34.0527 | -6.9664
3.9006 | -34.053 | -6.9773
3.9067 | -34.0532 | -6.9883
3.9128 | -34.0535 | -6.9989
3.9189 | -34.0538 | -7.0092
3.925 | -34.0541 | -7.0199
3.9311 | -34.0543 | -7.0299
3.9372 | -34.0547 | -7.0407
3.9434 | -34.0549 | -7.0513
3.9495 | -34.0553 | -7.062
3.9556 | -34.0557 | -7.0723
3.9617 | -34.0559 | -7.0824
3.9678 | -34.0563 | -7.0927
3.9739 | -34.0566 | -7.1032
3.98 | -34.0568 | -7.1139
3.9861 | -34.0571 | -7.1243
3.9922 | -34.0573 | -7.1349
3.9983 | -34.0577 | -7.145
4.0044 | -34.058 | -7.155
4.0105 | -34.0583 | -7.1656
4.0166 | -34.0586 | -7.1762
4.0227 | -34.0589 | -7.1866
4.0288 | -34.0591 | -7.1977
4.0349 | -34.0594 | -7.2079
4.041 | -34.0596 | -7.2183
4.0471 | -34.0599 | -7.2289
4.0532 | -34.0602 | -7.2391
4.0593 | -34.0605 | -7.2498
4.0654 | -34.0608 | -7.26
4.0715 | -34.0611 | -7.271
4.0776 | -34.0614 | -7.2815
4.0838 | -34.0617 | -7.2924
4.0899 | -34.062 | -7.3026
4.096 | -34.0623 | -7.313
4.1021 | -34.0626 | -7.3233
4.1082 | -34.0628 | -7.3335
4.1143 | -34.0631 | -7.3436
4.1204 | -34.0634 | -7.3538
4.1265 | -34.0637 | -7.3639
4.1326 | -34.064 | -7.3739
4.1387 | -34.0643 | -7.3842
4.1448 | -34.0646 | -7.3947
4.1509 | -34.0648 | -7.4051
4.157 | -34.065 | -7.416
4.1631 | -34.0653 | -7.4267
4.1692 | -34.0656 | -7.4369
4.1753 | -34.0659 | -7.4472
4.1814 | -34.0661 | -7.4572
4.1875 | -34.0664 | -7.4679
4.1936 | -34.0667 | -7.4784
4.1997 | -34.0669 | -7.4884
4.2058 | -34.0672 | -7.4991
4.2119 | -34.0674 | -7.5093
4.218 | -34.0676 | -7.5194
4.2241 | -34.0679 | -7.5295
4.2303 | -34.0682 | -7.5397
4.2364 | -34.0684 | -7.55
4.2425 | -34.0686 | -7.5603
4.2486 | -34.0688 | -7.5708
4.2547 | -34.069 | -7.5809
4.2608 | -34.0692 | -7.5913
4.2669 | -34.0694 | -7.6012
4.273 | -34.0696 | -7.6117
4.2791 | -34.0699 | -7.622
4.2852 | -34.0701 | -7.6325
4.2913 | -34.0702 | -7.6424
4.2974 | -34.0705 | -7.653
4.3035 | -34.0707 | -7.6636
4.3096 | -34.0709 | -7.6741
4.3157 | -34.0711 | -7.6841
4.3218 | -34.0714 | -7.6943
4.3279 | -34.0716 | -7.7046
4.334 | -34.0718 | -7.7148
4.3401 | -34.072 | -7.7246
4.3462 | -34.0723 | -7.7349
4.3523 | -34.0725 | -7.7447
4.3584 | -34.0727 | -7.7545
4.3645 | -34.0729 | -7.7646
4.3707 | -34.0731 | -7.7745
4.3768 | -34.0733 | -7.7846
4.3829 | -34.0735 | -7.7952
4.389 | -34.0737 | -7.8054
4.3951 | -34.0739 | -7.8153
4.4012 | -34.0741 | -7.8259
4.4073 | -34.0743 | -7.8358
4.4134 | -34.0745 | -7.8459
4.4195 | -34.0748 | -7.856
4.4256 | -34.0751 | -7.8659
4.4317 | -34.0753 | -7.876
4.4378 | -34.0754 | -7.8858
4.4439 | -34.0756 | -7.8956
4.45 | -34.0758 | -7.9057
4.4561 | -34.0759 | -7.9158
4.4622 | -34.0761 | -7.9259
4.4683 | -34.0763 | -7.936
4.4744 | -34.0765 | -7.9464
4.4805 | -34.0767 | -7.9561
4.4866 | -34.0769 | -7.9661
4.4927 | -34.0771 | -7.9765
4.4988 | -34.0773 | -7.9865
4.5049 | -34.0775 | -7.9962
4.511 | -34.0778 | -8.0059
4.5172 | -34.078 | -8.016
4.5233 | -34.0781 | -8.0264
4.5294 | -34.0784 | -8.0363
4.5355 | -34.0786 | -8.0464
4.5416 | -34.0788 | -8.0564
4.5477 | -34.079 | -8.0665
4.5538 | -34.0792 | -8.076
4.5599 | -34.0794 | -8.086
4.566 | -34.0796 | -8.096
4.5721 | -34.0798 | -8.1061
4.5782 | -34.08 | -8.1164
4.5843 | -34.0802 | -8.1266
4.5904 | -34.0805 | -8.1366
4.5965 | -34.0807 | -8.1466
4.6026 | -34.0809 | -8.1571
4.6087 | -34.0811 | -8.167
4.6148 | -34.0813 | -8.1774
4.6209 | -34.0815 | -8.1874
4.627 | -34.0818 | -8.1977
4.6331 | -34.0819 | -8.2076
4.6392 | -34.0821 | -8.2177
4.6453 | -34.0823 | -8.2277
4.6514 | -34.0826 | -8.2379
4.6576 | -34.0828 | -8.2483
4.6637 | -34.0829 | -8.258
4.6698 | -34.0831 | -8.2679
4.6759 | -34.0833 | -8.2784
4.682 | -34.0835 | -8.2883
4.6881 | -34.0836 | -8.2983
4.6942 | -34.0839 | -8.3083
4.7003 | -34.084 | -8.3185
4.7064 | -34.0843 | -8.3283
4.7125 | -34.0845 | -8.3387
4.7186 | -34.0847 | -8.3484
4.7247 | -34.0849 | -8.3586
4.7308 | -34.0852 | -8.3686
4.7369 | -34.0854 | -8.3788
4.743 | -34.0856 | -8.3892
4.7491 | -34.0858 | -8.3997
4.7552 | -34.086 | -8.4099
4.7613 | -34.0861 | -8.4201
4.7674 | -34.0863 | -8.43
4.7735 | -34.0865 | -8.4406
4.7796 | -34.0868 | -8.451
4.7857 | -34.087 | -8.4612
4.7918 | -34.0872 | -8.4712
4.7979 | -34.0875 | -8.4812
4.8041 | -34.0877 | -8.4913
4.8102 | -34.0879 | -8.5011
4.8163 | -34.0881 | -8.5112
4.8224 | -34.0883 | -8.5213
4.8285 | -34.0886 | -8.5311
4.8346 | -34.0888 | -8.541
4.8407 | -34.089 | -8.5512
4.8468 | -34.0892 | -8.5617
4.8529 | -34.0895 | -8.5721
4.859 | -34.0897 | -8.5826
4.8651 | -34.0899 | -8.5925
4.8712 | -34.0901 | -8.6024
4.8773 | -34.0903 | -8.6126
4.8834 | -34.0906 | -8.6228
4.8895 | -34.0908 | -8.6327
4.8956 | -34.091 | -8.6428
4.9017 | -34.0911 | -8.6527
4.9078 | -34.0914 | -8.6632
4.9139 | -34.0916 | -8.6729
4.92 | -34.0919 | -8.6828
4.9261 | -34.0921 | -8.6925
4.9322 | -34.0924 | -8.7027
4.9383 | -34.0926 | -8.7128
4.9445 | -34.0929 | -8.7229
4.9506 | -34.0931 | -8.7329
4.9567 | -34.0933 | -8.7428
4.9628 | -34.0935 | -8.7528
4.9689 | -34.0937 | -8.7633
4.975 | -34.0939 | -8.7738
4.9811 | -34.0942 | -8.7836
4.9872 | -34.0945 | -8.7936
4.9933 | -34.0948 | -8.8041
4.9994 | -34.095 | -8.8146
5.0055 | -34.0953 | -8.8248
5.0116 | -34.0956 | -8.8354
5.0177 | -34.0959 | -8.8459
5.0238 | -34.0961 | -8.8564
5.0299 | -34.0964 | -8.867
5.036 | -34.0967 | -8.8772
5.0421 | -34.0969 | -8.8872
5.0482 | -34.0972 | -8.8974
5.0543 | -34.0974 | -8.9073
5.0604 | -34.0977 | -8.9175
5.0665 | -34.098 | -8.9274
5.0726 | -34.0983 | -8.938
5.0787 | -34.0985 | -8.9484
5.0848 | -34.0988 | -8.9587
5.091 | -34.099 | -8.9686
5.0971 | -34.0993 | -8.9785
5.1032 | -34.0996 | -8.9891
5.1093 | -34.0999 | -8.9993
5.1154 | -34.1001 | -9.0092
5.1215 | -34.1003 | -9.0192
5.1276 | -34.1006 | -9.0291
5.1337 | -34.1009 | -9.039
5.1398 | -34.1012 | -9.0489
5.1459 | -34.1015 | -9.0594
5.152 | -34.1017 | -9.0698
5.1581 | -34.102 | -9.0802
5.1642 | -34.1023 | -9.0905
5.1703 | -34.1025 | -9.1007
5.1764 | -34.1029 | -9.1107
5.1825 | -34.1031 | -9.1207
5.1886 | -34.1034 | -9.1305
5.1947 | -34.1036 | -9.1408
5.2008 | -34.104 | -9.1511
5.2069 | -34.1042 | -9.1616
5.213 | -34.1045 | -9.1715
5.2191 | -34.1048 | -9.1818
5.2252 | -34.1051 | -9.1921
5.2314 | -34.1053 | -9.2025
5.2375 | -34.1056 | -9.2124
5.2436 | -34.1058 | -9.2227
5.2497 | -34.1061 | -9.2326
5.2558 | -34.1064 | -9.2432
5.2619 | -34.1067 | -9.2537
5.268 | -34.107 | -9.2641
5.2741 | -34.1074 | -9.2745
5.2802 | -34.1076 | -9.2851
5.2863 | -34.1079 | -9.2952
5.2924 | -34.1082 | -9.3057
5.2985 | -34.1086 | -9.316
5.3046 | -34.1088 | -9.3264
5.3107 | -34.1091 | -9.3369
5.3168 | -34.1094 | -9.3472
5.3229 | -34.1097 | -9.3575
5.329 | -34.11 | -9.3678
5.3351 | -34.1103 | -9.3777
5.3412 | -34.1106 | -9.3879
5.3473 | -34.1109 | -9.3983
5.3534 | -34.1112 | -9.4087
5.3595 | -34.1115 | -9.4191
5.3656 | -34.1118 | -9.4293
5.3717 | -34.1121 | -9.44
5.3779 | -34.1124 | -9.4507
5.384 | -34.1126 | -9.461
5.3901 | -34.113 | -9.4716
5.3962 | -34.1132 | -9.4816
5.4023 | -34.1136 | -9.4922
5.4084 | -34.1139 | -9.5025
5.4145 | -34.1142 | -9.5125
5.4206 | -34.1145 | -9.5228
5.4267 | -34.1148 | -9.5335
5.4328 | -34.1151 | -9.5439
5.4389 | -34.1154 | -9.5539
5.445 | -34.1157 | -9.5646
5.4511 | -34.116 | -9.5747
5.4572 | -34.1163 | -9.5852
5.4633 | -34.1166 | -9.5955
5.4694 | -34.1169 | -9.6058
5.4755 | -34.1172 | -9.6159
5.4816 | -34.1175 | -9.6262
5.4877 | -34.1178 | -9.6365
5.4938 | -34.1181 | -9.6465
5.4999 | -34.1184 | -9.6567
5.506 | -34.1188 | -9.667
5.5121 | -34.1191 | -9.6772
5.5183 | -34.1194 | -9.6873
5.5244 | -34.1196 | -9.6979
5.5305 | -34.12 | -9.7081
5.5366 | -34.1202 | -9.7181
5.5427 | -34.1205 | -9.7286
5.5488 | -34.1208 | -9.739
5.5549 | -34.1211 | -9.7493
5.561 | -34.1214 | -9.7599
5.5671 | -34.1217 | -9.7701
5.5732 | -34.122 | -9.7806
5.5793 | -34.1224 | -9.7914
5.5854 | -34.1227 | -9.8016
5.5915 | -34.123 | -9.8123
5.5976 | -34.1233 | -9.8227
5.6037 | -34.1236 | -9.8331
5.6098 | -34.1238 | -9.8434
5.6159 | -34.1242 | -9.8535
5.622 | -34.1245 | -9.864
5.6281 | -34.1248 | -9.8744
5.6342 | -34.1251 | -9.8851
5.6403 | -34.1254 | -9.8955
5.6464 | -34.1257 | -9.906
5.6525 | -34.126 | -9.9164
5.6586 | -34.1263 | -9.9265
5.6648 | -34.1267 | -9.9369
5.6709 | -34.127 | -9.9474
5.677 | -34.1273 | -9.9578
5.6831 | -34.1276 | -9.9682
5.6892 | -34.1279 | -9.9788
5.6953 | -34.1282 | -9.9891
5.7014 | -34.1285 | -9.9996
5.7075 | -34.1288 | -10.0099
5.7136 | -34.1291 | -10.0202
5.7197 | -34.1294 | -10.0306
5.7258 | -34.1297 | -10.0407
5.7319 | -34.13 | -10.0511
5.738 | -34.1303 | -10.0615
5.7441 | -34.1307 | -10.0718
5.7502 | -34.131 | -10.0818
5.7563 | -34.1313 | -10.0922
5.7624 | -34.1316 | -10.1028
5.7685 | -34.132 | -10.1132
5.7746 | -34.1322 | -10.1238
5.7807 | -34.1325 | -10.1342
5.7868 | -34.1328 | -10.1448
5.7929 | -34.133 | -10.1551
5.799 | -34.1334 | -10.1653
5.8052 | -34.1336 | -10.1754
5.8113 | -34.134 | -10.1857
5.8174 | -34.1342 | -10.1959
5.8235 | -34.1346 | -10.2062
5.8296 | -34.1348 | -10.2167
5.8357 | -34.1351 | -10.227
5.8418 | -34.1354 | -10.2375
5.8479 | -34.1357 | -10.2479
5.854 | -34.136 | -10.2584
5.8601 | -34.1363 | -10.2686
5.8662 | -34.1366 | -10.2791
5.8723 | -34.1369 | -10.2893
5.8784 | -34.1372 | -10.2995
5.8845 | -34.1375 | -10.3097
5.8906 | -34.1378 | -10.3201
5.8967 | -34.138 | -10.3303
5.9028 | -34.1384 | -10.3404
5.9089 | -34.1388 | -10.3506
5.915 | -34.139 | -10.3607
5.9211 | -34.1393 | -10.371
5.9272 | -34.1396 | -10.381
5.9333 | -34.1399 | -10.3912
5.9394 | -34.1402 | -10.4015
5.9455 | -34.1406 | -10.4115
5.9517 | -34.1409 | -10.4216
5.9578 | -34.1413 | -10.4321
5.9639 | -34.1416 | -10.442
5.97 | -34.1419 | -10.4523
5.9761 | -34.1422 | -10.4625
5.9822 | -34.1425 | -10.4729
5.9883 | -34.1428 | -10.4831
5.9944 | -34.1432 | -10.4935
6.0005 | -34.1435 | -10.5039
6.0066 | -34.1438 | -10.514
6.0127 | -34.1441 | -10.524
6.0188 | -34.1444 | -10.5345
6.0249 | -34.1447 | -10.5449
6.031 | -34.145 | -10.555
6.0371 | -34.1453 | -10.5652
6.0432 | -34.1456 | -10.5753
6.0493 | -34.1459 | -10.5854
6.0554 | -34.1463 | -10.5955
6.0615 | -34.1466 | -10.6057
6.0676 | -34.1469 | -10.6162
6.0737 | -34.1473 | -10.6265
6.0798 | -34.1475 | -10.6367
6.0859 | -34.1479 | -10.6469
6.0921 | -34.1482 | -10.6569
6.0982 | -34.1485 | -10.6671
6.1043 | -34.1488 | -10.6772
6.1104 | -34.1492 | -10.6876
6.1165 | -34.1495 | -10.6974
6.1226 | -34.1498 | -10.7075
6.1287 | -34.1501 | -10.7176
6.1348 | -34.1505 | -10.7278
6.1409 | -34.1508 | -10.7379
6.147 | -34.1511 | -10.748
6.1531 | -34.1514 | -10.7582
6.1592 | -34.1517 | -10.7682
6.1653 | -34.152 | -10.7782
6.1714 | -34.1523 | -10.7884
6.1775 | -34.1526 | -10.7987
6.1836 | -34.1529 | -10.8089
6.1897 | -34.1532 | -10.8191
6.1958 | -34.1535 | -10.8295
6.2019 | -34.1539 | -10.8398
6.208 | -34.1542 | -10.85
6.2141 | -34.1545 | -10.8601
6.2202 | -34.1549 | -10.8704
6.2263 | -34.1552 | -10.8804
6.2325 | -34.1555 | -10.8902
6.2386 | -34.1558 | -10.9
6.2447 | -34.1561 | -10.9105
6.2508 | -34.1564 | -10.9206
6.2569 | -34.1568 | -10.9307
6.263 | -34.1571 | -10.9408
6.2691 | -34.1574 | -10.9508
6.2752 | -34.1577 | -10.9611
6.2813 | -34.158 | -10.9711
6.2874 | -34.1583 | -10.9811
6.2935 | -34.1586 | -10.9912
6.2996 | -34.1589 | -11.0013
6.3057 | -34.1592 | -11.0115
6.3118 | -34.1595 | -11.0217
6.3179 | -34.1598 | -11.0318
6.324 | -34.1601 | -11.0418
6.3301 | -34.1604 | -11.0519
6.3362 | -34.1608 | -11.0619
6.3423 | -34.1611 | -11.0722
6.3484 | -34.1614 | -11.0824
6.3545 | -34.1617 | -11.0921
6.3606 | -34.162 | -11.1019
6.3667 | -34.1622 | -11.1118
6.3728 | -34.1626 | -11.122
6.379 | -34.1629 | -11.1321
6.3851 | -34.1632 | -11.1419
6.3912 | -34.1635 | -11.1519
6.3973 | -34.1638 | -11.162
6.4034 | -34.1641 | -11.1718
6.4095 | -34.1644 | -11.1818
6.4156 | -34.1647 | -11.1919
6.4217 | -34.165 | -11.2019
6.4278 | -34.1653 | -11.2121
6.4339 | -34.1656 | -11.2221
6.44 | -34.166 | -11.232
6.4461 | -34.1662 | -11.2423
6.4522 | -34.1666 | -11.2521
6.4583 | -34.1668 | -11.2623
6.4644 | -34.1671 | -11.2724
6.4705 | -34.1674 | -11.2821
6.4766 | -34.1677 | -11.2922
6.4827 | -34.168 | -11.3022
6.4888 | -34.1683 | -11.3121
6.4949 | -34.1686 | -11.3218
6.501 | -34.1689 | -11.3319
6.5071 | -34.1692 | -11.3416
6.5132 | -34.1696 | -11.3517
6.5194 | -34.1698 | -11.3617
6.5255 | -34.1702 | -11.3715
6.5316 | -34.1704 | -11.3816
6.5377 | -34.1707 | -11.3916
6.5438 | -34.171 | -11.4013
6.5499 | -34.1713 | -11.4112
6.556 | -34.1716 | -11.4212
6.5621 | -34.1719 | -11.4314
6.5682 | -34.1722 | -11.4414
6.5743 | -34.1726 | -11.4514
6.5804 | -34.1729 | -11.4614
6.5865 | -34.1732 | -11.4714
6.5926 | -34.1735 | -11.4813
6.5987 | -34.1738 | -11.491
6.6048 | -34.1741 | -11.501
6.6109 | -34.1744 | -11.5109
6.617 | -34.1747 | -11.521
6.6231 | -34.175 | -11.5309
6.6292 | -34.1753 | -11.5409
6.6353 | -34.1756 | -11.5511
6.6414 | -34.1759 | -11.561
6.6475 | -34.1762 | -11.5709
6.6536 | -34.1765 | -11.5808
6.6597 | -34.1768 | -11.5904
6.6659 | -34.1771 | -11.6004
6.672 | -34.1774 | -11.6102
6.6781 | -34.1777 | -11.6201
6.6842 | -34.1781 | -11.6299
6.6903 | -34.1784 | -11.6398
6.6964 | -34.1787 | -11.6498
6.7025 | -34.179 | -11.6595
6.7086 | -34.1793 | -11.6694
6.7147 | -34.1796 | -11.6794
6.7208 | -34.1799 | -11.6894
6.7269 | -34.1802 | -11.6993
6.733 | -34.1805 | -11.709
6.7391 | -34.1808 | -11.7188
6.7452 | -34.1811 | -11.7285
6.7513 | -34.1814 | -11.7381
6.7574 | -34.1817 | -11.7481
6.7635 | -34.182 | -11.7577
6.7696 | -34.1824 | -11.7675
6.7757 | -34.1827 | -11.7774
6.7818 | -34.183 | -11.7873
6.7879 | -34.1833 | -11.797
6.794 | -34.1836 | -11.8066
6.8001 | -34.1839 | -11.8164
6.8063 | -34.1842 | -11.8262
6.8124 | -34.1845 | -11.8363
6.8185 | -34.1848 | -11.8463
6.8246 | -34.1851 | -11.8559
6.8307 | -34.1854 | -11.8657
6.8368 | -34.1857 | -11.8758
6.8429 | -34.186 | -11.8856
6.849 | -34.1863 | -11.8955
6.8551 | -34.1867 | -11.9052
6.8612 | -34.187 | -11.915
6.8673 | -34.1873 | -11.9248
6.8734 | -34.1876 | -11.9347
6.8795 | -34.1879 | -11.9446
6.8856 | -34.1882 | -11.9544
6.8917 | -34.1885 | -11.9642
6.8978 | -34.1888 | -11.9737
6.9039 | -34.1891 | -11.9832
6.91 | -34.1895 | -11.9931
6.9161 | -34.1898 | -12.0026
6.9222 | -34.1901 | -12.0124
6.9283 | -34.1904 | -12.0219
6.9344 | -34.1907 | -12.0318
6.9405 | -34.1911 | -12.0415
6.9466 | -34.1914 | -12.0515
6.9528 | -34.1917 | -12.0615
6.9589 | -34.1921 | -12.0714
6.965 | -34.1924 | -12.0813
6.9711 | -34.1927 | -12.0912
6.9772 | -34.1931 | -12.101
6.9833 | -34.1934 | -12.1107
6.9894 | -34.1937 | -12.1208
6.9955 | -34.194 | -12.1303
7.0016 | -34.1944 | -12.14
7.0077 | -34.1947 | -12.1498
7.0138 | -34.195 | -12.1595
7.0199 | -34.1953 | -12.1691
7.026 | -34.1956 | -12.1788
7.0321 | -34.196 | -12.1884
7.0382 | -34.1963 | -12.1978
7.0443 | -34.1967 | -12.2077
7.0504 | -34.197 | -12.2176
7.0565 | -34.1973 | -12.2274
7.0626 | -34.1976 | -12.2368
7.0687 | -34.1979 | -12.2461
7.0748 | -34.1983 | -12.2556
7.0809 | -34.1986 | -12.2652
7.087 | -34.1989 | -12.275
7.0932 | -34.1992 | -12.2848
7.0993 | -34.1996 | -12.2945
7.1054 | -34.1999 | -12.304
7.1115 | -34.2003 | -12.3135
7.1176 | -34.2006 | -12.3231
7.1237 | -34.2009 | -12.3328
7.1298 | -34.2012 | -12.3424
7.1359 | -34.2015 | -12.3521
7.142 | -34.2018 | -12.3615
7.1481 | -34.2021 | -12.371
7.1542 | -34.2025 | -12.3808
7.1603 | -34.2028 | -12.3904
7.1664 | -34.2032 | -12.4
7.1725 | -34.2035 | -12.4098
7.1786 | -34.2039 | -12.4194
7.1847 | -34.2042 | -12.4291
7.1908 | -34.2045 | -12.439
7.1969 | -34.2048 | -12.4488
7.203 | -34.2051 | -12.4583
7.2091 | -34.2055 | -12.4679
7.2152 | -34.2058 | -12.4776
7.2213 | -34.2061 | -12.4871
7.2274 | -34.2065 | -12.4967
7.2335 | -34.2068 | -12.5062
7.2397 | -34.2071 | -12.5158
7.2458 | -34.2075 | -12.5252
7.2519 | -34.2079 | -12.5349
7.258 | -34.2082 | -12.5447
7.2641 | -34.2085 | -12.5543
7.2702 | -34.2088 | -12.5642
7.2763 | -34.2092 | -12.5739
7.2824 | -34.2095 | -12.5837
7.2885 | -34.2099 | -12.5934
7.2946 | -34.2102 | -12.603
7.3007 | -34.2105 | -12.6128
7.3068 | -34.2108 | -12.6227
7.3129 | -34.2111 | -12.6325
7.319 | -34.2115 | -12.6422
7.3251 | -34.2118 | -12.6516
7.3312 | -34.2122 | -12.6613
7.3373 | -34.2125 | -12.6707
7.3434 | -34.2129 | -12.6802
7.3495 | -34.2132 | -12.6897
7.3556 | -34.2136 | -12.6995
7.3617 | -34.2139 | -12.7093
7.3678 | -34.2142 | -12.7188
7.3739 | -34.2145 | -12.7287
7.3801 | -34.2149 | -12.7386
7.3862 | -34.2152 | -12.7482
7.3923 | -34.2156 | -12.758
7.3984 | -34.2159 | -12.7674
7.4045 | -34.2163 | -12.777
7.4106 | -34.2167 | -12.7866
7.4167 | -34.217 | -12.7963
7.4228 | -34.2173 | -12.8061
7.4289 | -34.2177 | -12.8157
7.435 | -34.218 | -12.8251
7.4411 | -34.2184 | -12.8349
7.4472 | -34.2187 | -12.8446
7.4533 | -34.219 | -12.854
7.4594 | -34.2194 | -12.8639
7.4655 | -34.2197 | -12.8734
7.4716 | -34.2201 | -12.8832
7.4777 | -34.2204 | -12.8928
7.4838 | -34.2208 | -12.9023
7.4899 | -34.2211 | -12.9119
7.496 | -34.2214 | -12.9214
7.5021 | -34.2218 | -12.931
7.5082 | -34.2221 | -12.9406
7.5143 | -34.2225 | -12.9503
7.5204 | -34.2228 | -12.9601
7.5266 | -34.2231 | -12.9696
7.5327 | -34.2235 | -12.9791
7.5388 | -34.2238 | -12.9891
7.5449 | -34.2242 | -12.999
7.551 | -34.2245 | -13.0085
7.5571 | -34.2249 | -13.0182
7.5632 | -34.2253 | -13.0279
7.5693 | -34.2256 | -13.0375
7.5754 | -34.226 | -13.047
7.5815 | -34.2263 | -13.0566
7.5876 | -34.2266 | -13.0665
7.5937 | -34.227 | -13.0764
7.5998 | -34.2274 | -13.086
7.6059 | -34.2278 | -13.0958
7.612 | -34.2281 | -13.1054
7.6181 | -34.2285 | -13.115
7.6242 | -34.2288 | -13.125
7.6303 | -34.2292 | -13.1349
7.6364 | -34.2295 | -13.1446
7.6425 | -34.2298 | -13.1543
7.6486 | -34.2301 | -13.164
7.6547 | -34.2305 | -13.1736
7.6608 | -34.2309 | -13.1832
7.667 | -34.2312 | -13.193
7.6731 | -34.2315 | -13.2028
7.6792 | -34.2319 | -13.2127
7.6853 | -34.2323 | -13.2225
7.6914 | -34.2326 | -13.2323
7.6975 | -34.233 | -13.2421
7.7036 | -34.2333 | -13.2519
7.7097 | -34.2337 | -13.2618
7.7158 | -34.2341 | -13.2716
7.7219 | -34.2344 | -13.2811
7.728 | -34.2348 | -13.291
7.7341 | -34.2351 | -13.3008
7.7402 | -34.2355 | -13.3106
7.7463 | -34.2358 | -13.3204
7.7524 | -34.2362 | -13.3304
7.7585 | -34.2365 | -13.3402
7.7646 | -34.2369 | -13.3499
7.7707 | -34.2372 | -13.3599
7.7768 | -34.2375 | -13.3698
7.7829 | -34.2379 | -13.3798
7.789 | -34.2382 | -13.3898
7.7951 | -34.2386 | -13.3995
7.8012 | -34.2389 | -13.4093
7.8073 | -34.2393 | -13.4193
7.8135 | -34.2397 | -13.4291
7.8196 | -34.24 | -13.4391
7.8257 | -34.2404 | -13.4491
7.8318 | -34.2407 | -13.459
7.8379 | -34.2411 | -13.4686
7.844 | -34.2414 | -13.4784
7.8501 | -34.2418 | -13.4883
7.8562 | -34.2421 | -13.4983
7.8623 | -34.2425 | -13.508
7.8684 | -34.2429 | -13.5179
7.8745 | -34.2432 | -13.5277
7.8806 | -34.2435 | -13.5378
7.8867 | -34.2439 | -13.5473
7.8928 | -34.2443 | -13.5569
7.8989 | -34.2446 | -13.5667
7.905 | -34.245 | -13.5766
7.9111 | -34.2453 | -13.5865
7.9172 | -34.2457 | -13.5963
7.9233 | -34.246 | -13.6062
7.9294 | -34.2464 | -13.6161
7.9355 | -34.2467 | -13.6257
7.9416 | -34.247 | -13.6356
7.9477 | -34.2474 | -13.6452
7.9539 | -34.2477 | -13.655
7.96 | -34.2481 | -13.6649
7.9661 | -34.2485 | -13.6748
7.9722 | -34.2489 | -13.6845
7.9783 | -34.2492 | -13.6946
7.9844 | -34.2496 | -13.7044
7.9905 | -34.2499 | -13.7145
7.9966 | -34.2503 | -13.7244
8.0027 | -34.2506 | -13.7343
8.0088 | -34.251 | -13.7441
8.0149 | -34.2513 | -13.7539
8.021 | -34.2516 | -13.7639
8.0271 | -34.252 | -13.774
8.0332 | -34.2524 | -13.7838
8.0393 | -34.2528 | -13.7938
8.0454 | -34.2531 | -13.8039
8.0515 | -34.2535 | -13.8138
8.0576 | -34.2538 | -13.8239
8.0637 | -34.2542 | -13.8337
8.0698 | -34.2545 | -13.8434
8.0759 | -34.2548 | -13.8531
8.082 | -34.2552 | -13.8628
8.0881 | -34.2555 | -13.8728
8.0942 | -34.2559 | -13.8827
8.1004 | -34.2563 | -13.8927
8.1065 | -34.2567 | -13.9024
8.1126 | -34.257 | -13.9123
8.1187 | -34.2573 | -13.9222
8.1248 | -34.2577 | -13.9319
8.1309 | -34.258 | -13.9416
8.137 | -34.2584 | -13.9512
8.1431 | -34.2588 | -13.9613
8.1492 | -34.2591 | -13.9712
8.1553 | -34.2595 | -13.9811
8.1614 | -34.2598 | -13.9912
8.1675 | -34.2601 | -14.001
8.1736 | -34.2605 | -14.0111
8.1797 | -34.2608 | -14.021
8.1858 | -34.2612 | -14.031
8.1919 | -34.2615 | -14.0408
8.198 | -34.2619 | -14.0505
8.2041 | -34.2623 | -14.0604
8.2102 | -34.2626 | -14.0703
8.2163 | -34.263 | -14.0799
8.2224 | -34.2633 | -14.0898
8.2285 | -34.2637 | -14.0994
8.2346 | -34.264 | -14.1092
8.2408 | -34.2644 | -14.1189
8.2469 | -34.2648 | -14.1288
8.253 | -34.2651 | -14.1386
8.2591 | -34.2655 | -14.1485
8.2652 | -34.2659 | -14.1584
8.2713 | -34.2662 | -14.1683
8.2774 | -34.2666 | -14.1783
8.2835 | -34.267 | -14.1884
8.2896 | -34.2673 | -14.1981
8.2957 | -34.2677 | -14.208
8.3018 | -34.268 | -14.2179
8.3079 | -34.2684 | -14.2276
8.314 | -34.2688 | -14.2374
8.3201 | -34.2692 | -14.2472
8.3262 | -34.2695 | -14.2571
8.3323 | -34.2699 | -14.2669
8.3384 | -34.2702 | -14.2769
8.3445 | -34.2706 | -14.2867
8.3506 | -34.2709 | -14.2966
8.3567 | -34.2713 | -14.3065
8.3628 | -34.2717 | -14.3164
8.3689 | -34.272 | -14.3261
8.375 | -34.2724 | -14.336
8.3812 | -34.2727 | -14.3458
8.3873 | -34.2731 | -14.3556
8.3934 | -34.2735 | -14.3656
8.3995 | -34.2738 | -14.3753
8.4056 | -34.2741 | -14.3851
8.4117 | -34.2745 | -14.395
8.4178 | -34.2748 | -14.4049
8.4239 | -34.2752 | -14.4146
8.43 | -34.2755 | -14.4244
8.4361 | -34.2759 | -14.4344
8.4422 | -34.2762 | -14.4441
8.4483 | -34.2766 | -14.4538
8.4544 | -34.277 | -14.4637
8.4605 | -34.2773 | -14.4734
8.4666 | -34.2777 | -14.4834
8.4727 | -34.278 | -14.4933
8.4788 | -34.2784 | -14.5031
8.4849 | -34.2787 | -14.5129
8.491 | -34.2791 | -14.5228
8.4971 | -34.2795 | -14.5323
8.5032 | -34.2798 | -14.542
8.5093 | -34.2802 | -14.5517
8.5154 | -34.2805 | -14.5616
8.5215 | -34.2809 | -14.5713
8.5277 | -34.2812 | -14.5811
8.5338 | -34.2816 | -14.5908
8.5399 | -34.2819 | -14.6006
8.546 | -34.2823 | -14.6102
8.5521 | -34.2827 | -14.62
8.5582 | -34.283 | -14.6298
8.5643 | -34.2834 | -14.6393
8.5704 | -34.2837 | -14.6488
8.5765 | -34.2841 | -14.6587
8.5826 | -34.2845 | -14.6681
8.5887 | -34.2849 | -14.6778
8.5948 | -34.2852 | -14.6877
8.6009 | -34.2856 | -14.6972
8.607 | -34.2859 | -14.7071
8.6131 | -34.2863 | -14.7169
8.6192 | -34.2867 | -14.7265
8.6253 | -34.2871 | -14.7362
8.6314 | -34.2874 | -14.7458
8.6375 | -34.2877 | -14.7556
8.6436 | -34.2881 | -14.7654
8.6497 | -34.2885 | -14.7749
8.6558 | -34.2888 | -14.7846
8.6619 | -34.2891 | -14.7943
8.6681 | -34.2895 | -14.8042
8.6742 | -34.2899 | -14.8139
8.6803 | -34.2902 | -14.8236
8.6864 | -34.2906 | -14.833
8.6925 | -34.2909 | -14.8426
8.6986 | -34.2913 | -14.8523
8.7047 | -34.2916 | -14.8622
8.7108 | -34.292 | -14.8719
8.7169 | -34.2923 | -14.8817
8.723 | -34.2927 | -14.8912
8.7291 | -34.2931 | -14.9007
8.7352 | -34.2934 | -14.9106
8.7413 | -34.2938 | -14.9203
8.7474 | -34.2941 | -14.9298
8.7535 | -34.2945 | -14.9395
8.7596 | -34.2949 | -14.9492
8.7657 | -34.2953 | -14.959
8.7718 | -34.2956 | -14.9687
8.7779 | -34.296 | -14.9785
8.784 | -34.2963 | -14.9882
8.7901 | -34.2967 | -14.9979
8.7962 | -34.2971 | -15.0073
8.8023 | -34.2974 | -15.0169
8.8084 | -34.2978 | -15.0265
8.8146 | -34.2982 | -15.0361
8.8207 | -34.2985 | -15.0457
8.8268 | -34.2989 | -15.055
8.8329 | -34.2992 | -15.0646
8.839 | -34.2996 | -15.0743
8.8451 | -34.3 | -15.0839
8.8512 | -34.3003 | -15.0934
8.8573 | -34.3007 | -15.1031
8.8634 | -34.3011 | -15.1125
8.8695 | -34.3014 | -15.1223
8.8756 | -34.3017 | -15.1319
8.8817 | -34.3021 | -15.1414
8.8878 | -34.3025 | -15.1513
8.8939 | -34.3028 | -15.1609
8.9 | -34.3032 | -15.1705
8.9061 | -34.3036 | -15.1801
8.9122 | -34.3039 | -15.1897
8.9183 | -34.3043 | -15.1994
8.9244 | -34.3047 | -15.2091
8.9305 | -34.3051 | -15.2185
8.9366 | -34.3055 | -15.2281
8.9427 | -34.3058 | -15.2378
8.9488 | -34.3062 | -15.2473
8.955 | -34.3065 | -15.2568
8.9611 | -34.3069 | -15.2666
8.9672 | -34.3074 | -15.2761
8.9733 | -34.3077 | -15.2855
8.9794 | -34.3081 | -15.2949
8.9855 | -34.3085 | -15.3047
8.9916 | -34.3088 | -15.3144
8.9977 | -34.3092 | -15.324
9.0038 | -34.3096 | -15.3334
9.0099 | -34.31 | -15.3431
9.016 | -34.3104 | -15.3527
9.0221 | -34.3107 | -15.3621
9.0282 | -34.3111 | -15.3715
9.0343 | -34.3115 | -15.3812
9.0404 | -34.3119 | -15.3907
9.0465 | -34.3123 | -15.4005
9.0526 | -34.3127 | -15.4099
9.0587 | -34.3131 | -15.4195
9.0648 | -34.3135 | -15.4289
9.0709 | -34.3138 | -15.4385
9.077 | -34.3143 | -15.4481
9.0831 | -34.3146 | -15.4578
9.0892 | -34.315 | -15.4674
9.0953 | -34.3154 | -15.477
9.1015 | -34.3158 | -15.4865
9.1076 | -34.3162 | -15.4959
9.1137 | -34.3166 | -15.5054
9.1198 | -34.317 | -15.515
9.1259 | -34.3173 | -15.5245
9.132 | -34.3177 | -15.5342
9.1381 | -34.3181 | -15.5436
9.1442 | -34.3185 | -15.5528
9.1503 | -34.3189 | -15.5625
9.1564 | -34.3193 | -15.5719
9.1625 | -34.3197 | -15.5815
9.1686 | -34.3201 | -15.591
9.1747 | -34.3205 | -15.6005
9.1808 | -34.3209 | -15.6099
9.1869 | -34.3213 | -15.6194
9.193 | -34.3217 | -15.629
9.1991 | -34.322 | -15.6386
9.2052 | -34.3224 | -15.6481
9.2113 | -34.3228 | -15.6576
9.2174 | -34.3233 | -15.6671
9.2235 | -34.3237 | -15.6766
9.2296 | -34.3241 | -15.686
9.2357 | -34.3245 | -15.6957
9.2419 | -34.3249 | -15.7052
9.248 | -34.3253 | -15.7147
9.2541 | -34.3257 | -15.7243
9.2602 | -34.3261 | -15.7339
9.2663 | -34.3265 | -15.7431
9.2724 | -34.3269 | -15.7526
9.2785 | -34.3273 | -15.762
9.2846 | -34.3277 | -15.7713
9.2907 | -34.3281 | -15.7808
9.2968 | -34.3285 | -15.7904
9.3029 | -34.3289 | -15.8
9.309 | -34.3293 | -15.8095
9.3151 | -34.3297 | -15.8189
9.3212 | -34.3301 | -15.8283
9.3273 | -34.3305 | -15.8377
9.3334 | -34.3309 | -15.8471
9.3395 | -34.3313 | -15.8564
9.3456 | -34.3317 | -15.8657
9.3517 | -34.3321 | -15.8752
9.3578 | -34.3325 | -15.8847
9.3639 | -34.3329 | -15.894
9.37 | -34.3333 | -15.9035
9.3761 | -34.3337 | -15.9131
9.3822 | -34.3341 | -15.9225
9.3884 | -34.3346 | -15.9317
9.3945 | -34.3349 | -15.9412
9.4006 | -34.3354 | -15.9507
9.4067 | -34.3358 | -15.9602
9.4128 | -34.3362 | -15.9698
9.4189 | -34.3366 | -15.9795
9.425 | -34.337 | -15.9888
9.4311 | -34.3374 | -15.9982
9.4372 | -34.3379 | -16.0076
9.4433 | -34.3383 | -16.0171
9.4494 | -34.3387 | -16.0263
9.4555 | -34.3391 | -16.0359
9.4616 | -34.3395 | -16.0453
9.4677 | -34.34 | -16.0547
9.4738 | -34.3404 | -16.0641
9.4799 | -34.3408 | -16.0736
9.486 | -34.3412 | -16.083
9.4921 | -34.3416 | -16.0925
9.4982 | -34.342 | -16.1016
9.5043 | -34.3424 | -16.111
9.5104 | -34.3428 | -16.1203
9.5165 | -34.3432 | -16.1297
9.5226 | -34.3436 | -16.1393
9.5288 | -34.3441 | -16.1485
9.5349 | -34.3445 | -16.1579
9.541 | -34.3449 | -16.1672
9.5471 | -34.3453 | -16.1766
9.5532 | -34.3457 | -16.1861
9.5593 | -34.3461 | -16.1956
9.5654 | -34.3465 | -16.2049
9.5715 | -34.3469 | -16.2143
9.5776 | -34.3473 | -16.2238
9.5837 | -34.3477 | -16.2333
9.5898 | -34.3481 | -16.2428
9.5959 | -34.3485 | -16.2522
9.602 | -34.3489 | -16.2615
9.6081 | -34.3494 | -16.2708
9.6142 | -34.3498 | -16.2801
9.6203 | -34.3502 | -16.2895
9.6264 | -34.3506 | -16.2988
9.6325 | -34.351 | -16.3083
9.6386 | -34.3514 | -16.3175
9.6447 | -34.3518 | -16.327
9.6508 | -34.3522 | -16.3363
9.6569 | -34.3526 | -16.3457
9.663 | -34.3531 | -16.3551
9.6691 | -34.3535 | -16.3646
9.6753 | -34.3539 | -16.3739
9.6814 | -34.3543 | -16.3832
9.6875 | -34.3547 | -16.3924
9.6936 | -34.3551 | -16.4016
9.6997 | -34.3556 | -16.4111
9.7058 | -34.356 | -16.4204
9.7119 | -34.3565 | -16.4299
9.718 | -34.3569 | -16.4394
9.7241 | -34.3573 | -16.4487
9.7302 | -34.3577 | -16.4582
9.7363 | -34.3581 | -16.4674
9.7424 | -34.3585 | -16.4766
9.7485 | -34.3589 | -16.4861
9.7546 | -34.3593 | -16.4953
9.7607 | -34.3597 | -16.5049
9.7668 | -34.3602 | -16.5143
9.7729 | -34.3606 | -16.5236
9.779 | -34.3611 | -16.5328
9.7851 | -34.3615 | -16.5424
9.7912 | -34.3619 | -16.5516
9.7973 | -34.3624 | -16.5612
9.8034 | -34.3628 | -16.5704
9.8095 | -34.3632 | -16.58
9.8157 | -34.3637 | -16.5893
9.8218 | -34.3641 | -16.5989
9.8279 | -34.3645 | -16.6081
9.834 | -34.3648 | -16.6175
9.8401 | -34.3653 | -16.6269
9.8462 | -34.3657 | -16.636
9.8523 | -34.3661 | -16.6453
9.8584 | -34.3665 | -16.6547
9.8645 | -34.3669 | -16.6641
9.8706 | -34.3673 | -16.6735
9.8767 | -34.3677 | -16.6828
9.8828 | -34.3681 | -16.692
9.8889 | -34.3686 | -16.7012
9.895 | -34.369 | -16.7106
9.9011 | -34.3694 | -16.7199
9.9072 | -34.3698 | -16.7292
9.9133 | -34.3702 | -16.7385
9.9194 | -34.3706 | -16.7477
9.9255 | -34.3711 | -16.757
9.9316 | -34.3715 | -16.7663
9.9377 | -34.3719 | -16.7756
9.9438 | -34.3723 | -16.785
9.9499 | -34.3727 | -16.7946
9.956 | -34.3731 | -16.8039
9.9622 | -34.3736 | -16.8133
9.9683 | -34.374 | -16.8226
9.9744 | -34.3744 | -16.8321
9.9805 | -34.3748 | -16.8414
9.9866 | -34.3752 | -16.8506
9.9927 | -34.3756 | -16.8599
9.9988 | -34.376 | -16.8693
10.0049 | -34.3765 | -16.8786
10.011 | -34.3769 | -16.8881
10.0171 | -34.3773 | -16.8973
10.0232 | -34.3777 | -16.9066
10.0293 | -34.3782 | -16.916
10.0354 | -34.3786 | -16.9253
10.0415 | -34.3791 | -16.9346
10.0476 | -34.3795 | -16.9441
10.0537 | -34.3799 | -16.9535
10.0598 | -34.3803 | -16.9629
10.0659 | -34.3807 | -16.9723
10.072 | -34.3811 | -16.9813
10.0781 | -34.3815 | -16.9906
10.0842 | -34.382 | -16.9999
10.0903 | -34.3824 | -17.0092
10.0964 | -34.3828 | -17.0182
10.1026 | -34.3832 | -17.0276
10.1087 | -34.3836 | -17.0369
10.1148 | -34.384 | -17.0463
10.1209 | -34.3844 | -17.0554
10.127 | -34.3849 | -17.0646
10.1331 | -34.3853 | -17.0739
10.1392 | -34.3857 | -17.0829
10.1453 | -34.3861 | -17.0922
10.1514 | -34.3865 | -17.1016
10.1575 | -34.3869 | -17.111
10.1636 | -34.3874 | -17.12
10.1697 | -34.3878 | -17.1294
10.1758 | -34.3882 | -17.1385
10.1819 | -34.3886 | -17.1478
10.188 | -34.389 | -17.1571
10.1941 | -34.3894 | -17.1665
10.2002 | -34.3898 | -17.1758
10.2063 | -34.3903 | -17.185
10.2124 | -34.3907 | -17.1944
10.2185 | -34.3911 | -17.2038
10.2246 | -34.3915 | -17.213
10.2307 | -34.3919 | -17.2223
10.2368 | -34.3924 | -17.2313
10.2429 | -34.3928 | -17.2406
10.2491 | -34.3932 | -17.2499
10.2552 | -34.3936 | -17.2592
10.2613 | -34.394 | -17.2685
10.2674 | -34.3944 | -17.2779
10.2735 | -34.3948 | -17.2873
10.2796 | -34.3953 | -17.2966
10.2857 | -34.3957 | -17.306
10.2918 | -34.3961 | -17.3154
10.2979 | -34.3965 | -17.3247
10.304 | -34.397 | -17.3337
10.3101 | -34.3973 | -17.343
10.3162 | -34.3978 | -17.3525
10.3223 | -34.3982 | -17.3615
10.3284 | -34.3986 | -17.3708
10.3345 | -34.399 | -17.3799
10.3406 | -34.3994 | -17.3892
10.3467 | -34.3998 | -17.3984
10.3528 | -34.4002 | -17.4075
10.3589 | -34.4006 | -17.4168
10.365 | -34.401 | -17.4262
10.3711 | -34.4014 | -17.4356
10.3772 | -34.4019 | -17.4448
10.3833 | -34.4023 | -17.4541
10.3895 | -34.4027 | -17.4633
10.3956 | -34.4032 | -17.4726
10.4017 | -34.4036 | -17.4817
10.4078 | -34.404 | -17.4909
10.4139 | -34.4044 | -17.5003
10.42 | -34.4048 | -17.5097
10.4261 | -34.4052 | -17.5189
10.4322 | -34.4056 | -17.5281
10.4383 | -34.406 | -17.5374
10.4444 | -34.4064 | -17.5465
10.4505 | -34.4068 | -17.5558
10.4566 | -34.4072 | -17.5649
10.4627 | -34.4076 | -17.574
10.4688 | -34.4081 | -17.5831
10.4749 | -34.4085 | -17.5924
10.481 | -34.4089 | -17.6017
10.4871 | -34.4093 | -17.6108
10.4932 | -34.4097 | -17.6199
10.4993 | -34.4102 | -17.6293
10.5054 | -34.4106 | -17.6386
10.5115 | -34.411 | -17.6478
10.5176 | -34.4114 | -17.6569
10.5237 | -34.4118 | -17.6664
10.5298 | -34.4122 | -17.6754
10.536 | -34.4126 | -17.6847
10.5421 | -34.4131 | -17.6939
10.5482 | -34.4135 | -17.7032
10.5543 | -34.4139 | -17.7124
10.5604 | -34.4143 | -17.7214
10.5665 | -34.4147 | -17.7305
10.5726 | -34.4151 | -17.7398
10.5787 | -34.4155 | -17.749
10.5848 | -34.4159 | -17.758
10.5909 | -34.4164 | -17.7673
10.597 | -34.4168 | -17.7766
10.6031 | -34.4172 | -17.786
10.6092 | -34.4176 | -17.7952
10.6153 | -34.418 | -17.8046
10.6214 | -34.4184 | -17.8137
10.6275 | -34.4188 | -17.8229
10.6336 | -34.4192 | -17.8321
10.6397 | -34.4196 | -17.8413
10.6458 | -34.42 | -17.8505
10.6519 | -34.4205 | -17.8598
10.658 | -34.4209 | -17.869
10.6641 | -34.4213 | -17.8782
10.6702 | -34.4217 | -17.8875
10.6764 | -34.4221 | -17.8968
10.6825 | -34.4225 | -17.9059
10.6886 | -34.4229 | -17.9149
10.6947 | -34.4233 | -17.9241
10.7008 | -34.4238 | -17.9333
10.7069 | -34.4242 | -17.9426
10.713 | -34.4246 | -17.9518
10.7191 | -34.425 | -17.961
10.7252 | -34.4254 | -17.9699
10.7313 | -34.4258 | -17.9791
10.7374 | -34.4262 | -17.9882
10.7435 | -34.4266 | -17.9975
10.7496 | -34.427 | -18.0064
10.7557 | -34.4274 | -18.0157
10.7618 | -34.4279 | -18.0247
10.7679 | -34.4283 | -18.0339
10.774 | -34.4287 | -18.0431
10.7801 | -34.4291 | -18.0524
10.7862 | -34.4295 | -18.0613
10.7923 | -34.4299 | -18.0706
10.7984 | -34.4303 | -18.0797
10.8045 | -34.4307 | -18.0887
10.8106 | -34.4311 | -18.098
10.8168 | -34.4315 | -18.1071
10.8229 | -34.4319 | -18.1163
10.829 | -34.4323 | -18.1253
10.8351 | -34.4327 | -18.1343
10.8412 | -34.4332 | -18.1434
10.8473 | -34.4336 | -18.1523
10.8534 | -34.434 | -18.1615
10.8595 | -34.4344 | -18.1707
10.8656 | -34.4349 | -18.1797
10.8717 | -34.4353 | -18.1889
10.8778 | -34.4357 | -18.198
10.8839 | -34.4361 | -18.2072
10.89 | -34.4365 | -18.2164
10.8961 | -34.437 | -18.2255
10.9022 | -34.4374 | -18.2345
10.9083 | -34.4378 | -18.2436
10.9144 | -34.4382 | -18.2526
10.9205 | -34.4386 | -18.2617
10.9266 | -34.439 | -18.2708
10.9327 | -34.4394 | -18.2799
10.9388 | -34.4399 | -18.2889
10.9449 | -34.4403 | -18.298
10.951 | -34.4407 | -18.3071
10.9571 | -34.4411 | -18.3162
10.9633 | -34.4415 | -18.3255
10.9694 | -34.4419 | -18.3348
10.9755 | -34.4423 | -18.3439
10.9816 | -34.4427 | -18.3531
10.9877 | -34.4432 | -18.3621
10.9938 | -34.4436 | -18.3714
10.9999 | -34.444 | -18.3803
11.006 | -34.4444 | -18.3893
11.0121 | -34.4448 | -18.3984
11.0182 | -34.4453 | -18.4075
11.0243 | -34.4457 | -18.4166
11.0304 | -34.4461 | -18.4257
11.0365 | -34.4465 | -18.4347
11.0426 | -34.4469 | -18.4438
11.0487 | -34.4474 | -18.453
11.0548 | -34.4478 | -18.4618
11.0609 | -34.4482 | -18.4709
11.067 | -34.4486 | -18.48
11.0731 | -34.4491 | -18.489
11.0792 | -34.4495 | -18.4982
11.0853 | -34.4499 | -18.5072
11.0914 | -34.4503 | -18.5165
11.0975 | -34.4508 | -18.5256
11.1037 | -34.4512 | -18.5346
11.1098 | -34.4516 | -18.544
11.1159 | -34.452 | -18.5528
11.122 | -34.4524 | -18.5619
11.1281 | -34.4528 | -18.5711
11.1342 | -34.4533 | -18.5801
11.1403 | -34.4537 | -18.5894
11.1464 | -34.4542 | -18.5985
11.1525 | -34.4546 | -18.6074
11.1586 | -34.455 | -18.6166
11.1647 | -34.4554 | -18.6256
11.1708 | -34.4558 | -18.6347
11.1769 | -34.4562 | -18.6438
11.183 | -34.4566 | -18.6525
11.1891 | -34.4571 | -18.6616
11.1952 | -34.4575 | -18.6707
11.2013 | -34.4579 | -18.6799
11.2074 | -34.4583 | -18.689
11.2135 | -34.4588 | -18.6983
11.2196 | -34.4592 | -18.7072
11.2257 | -34.4597 | -18.7164
11.2318 | -34.4601 | -18.7252
11.2379 | -34.4605 | -18.7344
11.244 | -34.461 | -18.7435
11.2502 | -34.4614 | -18.7527
11.2563 | -34.4619 | -18.7617
11.2624 | -34.4623 | -18.7707
11.2685 | -34.4627 | -18.7797
11.2746 | -34.4631 | -18.7886
11.2807 | -34.4636 | -18.7979
11.2868 | -34.464 | -18.8071
11.2929 | -34.4645 | -18.8158
11.299 | -34.4649 | -18.8249
11.3051 | -34.4653 | -18.8341
11.3112 | -34.4658 | -18.8432
11.3173 | -34.4662 | -18.852
11.3234 | -34.4666 | -18.8609
11.3295 | -34.4671 | -18.8701
11.3356 | -34.4675 | -18.8791
11.3417 | -34.4679 | -18.888
11.3478 | -34.4684 | -18.8973
11.3539 | -34.4687 | -18.9063
11.36 | -34.4692 | -18.9153
11.3661 | -34.4696 | -18.9241
11.3722 | -34.4701 | -18.9328
11.3783 | -34.4705 | -18.942
11.3844 | -34.4709 | -18.951
11.3906 | -34.4714 | -18.96
11.3967 | -34.4718 | -18.9689
11.4028 | -34.4722 | -18.978
11.4089 | -34.4726 | -18.9868
11.415 | -34.473 | -18.9959
11.4211 | -34.4734 | -19.005
11.4272 | -34.4739 | -19.0141
11.4333 | -34.4743 | -19.0228
11.4394 | -34.4747 | -19.0318
11.4455 | -34.4751 | -19.0406
11.4516 | -34.4755 | -19.0496
11.4577 | -34.476 | -19.0586
11.4638 | -34.4764 | -19.0674
11.4699 | -34.4769 | -19.0765
11.476 | -34.4773 | -19.0858
11.4821 | -34.4777 | -19.0948
11.4882 | -34.4782 | -19.1039
11.4943 | -34.4786 | -19.113
11.5004 | -34.4791 | -19.1216
11.5065 | -34.4795 | -19.1305
11.5126 | -34.48 | -19.1394
11.5187 | -34.4804 | -19.1485
11.5248 | -34.4808 | -19.1577
11.5309 | -34.4813 | -19.1665
11.5371 | -34.4817 | -19.1754
11.5432 | -34.4821 | -19.1844
11.5493 | -34.4826 | -19.1935
11.5554 | -34.483 | -19.2025
11.5615 | -34.4835 | -19.2114
11.5676 | -34.4839 | -19.2205
11.5737 | -34.4844 | -19.2295
11.5798 | -34.4849 | -19.2382
11.5859 | -34.4853 | -19.2472
11.592 | -34.4857 | -19.2561
11.5981 | -34.4862 | -19.2649
11.6042 | -34.4866 | -19.2739
11.6103 | -34.487 | -19.2828
11.6164 | -34.4875 | -19.2916
11.6225 | -34.4879 | -19.3005
11.6286 | -34.4883 | -19.3095
11.6347 | -34.4888 | -19.3185
11.6408 | -34.4892 | -19.3273
11.6469 | -34.4897 | -19.3362
11.653 | -34.4901 | -19.3451
11.6591 | -34.4905 | -19.3541
11.6652 | -34.491 | -19.3629
11.6713 | -34.4914 | -19.3717
11.6775 | -34.4919 | -19.3804
11.6836 | -34.4923 | -19.3894
11.6897 | -34.4928 | -19.3983
11.6958 | -34.4932 | -19.4073
11.7019 | -34.4936 | -19.4163
11.708 | -34.4941 | -19.4251
11.7141 | -34.4946 | -19.4338
11.7202 | -34.495 | -19.4428
11.7263 | -34.4954 | -19.4517
11.7324 | -34.4959 | -19.4608
11.7385 | -34.4963 | -19.4698
11.7446 | -34.4967 | -19.4786
11.7507 | -34.4972 | -19.4873
11.7568 | -34.4976 | -19.4961
11.7629 | -34.4981 | -19.5052
11.769 | -34.4985 | -19.514
11.7751 | -34.4989 | -19.523
11.7812 | -34.4994 | -19.5316
11.7873 | -34.4999 | -19.5406
11.7934 | -34.5003 | -19.5495
11.7995 | -34.5007 | -19.5583
11.8056 | -34.5012 | -19.567
11.8117 | -34.5016 | -19.576
11.8178 | -34.5021 | -19.5848
11.824 | -34.5026 | -19.5937
11.8301 | -34.503 | -19.6025
11.8362 | -34.5035 | -19.6112
11.8423 | -34.5039 | -19.6201
11.8484 | -34.5044 | -19.6289
11.8545 | -34.5049 | -19.6377
11.8606 | -34.5053 | -19.6465
11.8667 | -34.5057 | -19.6555
11.8728 | -34.5062 | -19.6642
11.8789 | -34.5066 | -19.673
11.885 | -34.5071 | -19.6818
11.8911 | -34.5075 | -19.6905
11.8972 | -34.508 | -19.6992
11.9033 | -34.5084 | -19.7079
11.9094 | -34.5089 | -19.7167
11.9155 | -34.5093 | -19.7255
11.9216 | -34.5098 | -19.7342
11.9277 | -34.5103 | -19.7429
11.9338 | -34.5107 | -19.7517
11.9399 | -34.5111 | -19.7607
11.946 | -34.5116 | -19.7694
11.9521 | -34.512 | -19.7782
11.9582 | -34.5125 | -19.7869
11.9644 | -34.513 | -19.7956
11.9705 | -34.5134 | -19.8044
11.9766 | -34.5139 | -19.8132
11.9827 | -34.5144 | -19.822
11.9888 | -34.5149 | -19.8307
11.9949 | -34.5153 | -19.8392
12.001 | -34.5158 | -19.8481
12.0071 | -34.5162 | -19.8569
12.0132 | -34.5167 | -19.8655
12.0193 | -34.5171 | -19.8743
12.0254 | -34.5176 | -19.8829
12.0315 | -34.518 | -19.8917
12.0376 | -34.5185 | -19.9004
12.0437 | -34.5189 | -19.9091
12.0498 | -34.5194 | -19.918
12.0559 | -34.5198 | -19.9266
12.062 | -34.5203 | -19.9351
12.0681 | -34.5208 | -19.944
12.0742 | -34.5213 | -19.9526
12.0803 | -34.5217 | -19.9615
12.0864 | -34.5222 | -19.9703
12.0925 | -34.5226 | -19.979
12.0986 | -34.5231 | -19.9875
12.1047 | -34.5236 | -19.9961
12.1109 | -34.5241 | -20.0047
12.117 | -34.5245 | -20.0134
12.1231 | -34.525 | -20.022
12.1292 | -34.5254 | -20.0306
12.1353 | -34.5259 | -20.0392
12.1414 | -34.5264 | -20.048
12.1475 | -34.5268 | -20.0568
12.1536 | -34.5273 | -20.0655
12.1597 | -34.5277 | -20.0742
12.1658 | -34.5282 | -20.0827
12.1719 | -34.5287 | -20.0911
12.178 | -34.5291 | -20.0999
12.1841 | -34.5296 | -20.1087
12.1902 | -34.5301 | -20.1174
12.1963 | -34.5306 | -20.1261
12.2024 | -34.5311 | -20.1346
12.2085 | -34.5315 | -20.1433
12.2146 | -34.532 | -20.1521
12.2207 | -34.5324 | -20.1607
12.2268 | -34.5329 | -20.1693
12.2329 | -34.5334 | -20.1779
12.239 | -34.5339 | -20.1861
12.2451 | -34.5343 | -20.1949
12.2513 | -34.5348 | -20.2037
12.2574 | -34.5352 | -20.2124
12.2635 | -34.5357 | -20.221
12.2696 | -34.5362 | -20.2296
12.2757 | -34.5366 | -20.2379
12.2818 | -34.5371 | -20.2466
12.2879 | -34.5376 | -20.2553
12.294 | -34.5381 | -20.2638
12.3001 | -34.5385 | -20.2724
12.3062 | -34.539 | -20.2811
12.3123 | -34.5395 | -20.2894
12.3184 | -34.5399 | -20.2978
12.3245 | -34.5404 | -20.3063
12.3306 | -34.5409 | -20.3149
12.3367 | -34.5413 | -20.3235
12.3428 | -34.5418 | -20.332
12.3489 | -34.5423 | -20.3404
12.355 | -34.5427 | -20.3492
12.3611 | -34.5432 | -20.3579
12.3672 | -34.5437 | -20.3664
12.3733 | -34.5441 | -20.3749
12.3794 | -34.5446 | -20.383
12.3855 | -34.5451 | -20.3915
12.3916 | -34.5455 | -20.4
12.3978 | -34.546 | -20.4088
12.4039 | -34.5465 | -20.4174
12.41 | -34.547 | -20.4259
12.4161 | -34.5475 | -20.4342
12.4222 | -34.5479 | -20.4427
12.4283 | -34.5484 | -20.4512
12.4344 | -34.5489 | -20.4599
12.4405 | -34.5493 | -20.4685
12.4466 | -34.5498 | -20.4769
12.4527 | -34.5503 | -20.4854
12.4588 | -34.5507 | -20.494
12.4649 | -34.5512 | -20.5025
12.471 | -34.5517 | -20.5112
12.4771 | -34.5521 | -20.5197
12.4832 | -34.5526 | -20.5282
12.4893 | -34.5531 | -20.5367
12.4954 | -34.5536 | -20.545
12.5015 | -34.554 | -20.5535
12.5076 | -34.5545 | -20.5621
12.5137 | -34.555 | -20.5707
12.5198 | -34.5555 | -20.5792
12.5259 | -34.5559 | -20.5877
12.532 | -34.5564 | -20.5962
12.5382 | -34.5569 | -20.6048
12.5443 | -34.5573 | -20.6133
12.5504 | -34.5578 | -20.6218
12.5565 | -34.5583 | -20.63
12.5626 | -34.5588 | -20.6387
12.5687 | -34.5592 | -20.6472
12.5748 | -34.5597 | -20.6557
12.5809 | -34.5602 | -20.6642
12.587 | -34.5607 | -20.6728
12.5931 | -34.5612 | -20.6811
12.5992 | -34.5617 | -20.6896
12.6053 | -34.5621 | -20.6981
12.6114 | -34.5626 | -20.7065
12.6175 | -34.5631 | -20.7151
12.6236 | -34.5635 | -20.7236
12.6297 | -34.564 | -20.732
12.6358 | -34.5645 | -20.7402
12.6419 | -34.5649 | -20.7487
12.648 | -34.5655 | -20.7572
12.6541 | -34.5659 | -20.7656
12.6602 | -34.5664 | -20.7737
12.6663 | -34.5669 | -20.7822
12.6724 | -34.5674 | -20.7906
12.6785 | -34.5678 | -20.7989
12.6847 | -34.5683 | -20.8072
12.6908 | -34.5688 | -20.8156
12.6969 | -34.5693 | -20.8239
12.703 | -34.5697 | -20.8324
12.7091 | -34.5702 | -20.8408
12.7152 | -34.5707 | -20.8492
12.7213 | -34.5711 | -20.8578
12.7274 | -34.5716 | -20.8663
12.7335 | -34.5721 | -20.8742
12.7396 | -34.5726 | -20.8825
12.7457 | -34.573 | -20.8911
12.7518 | -34.5735 | -20.8994
12.7579 | -34.574 | -20.9079
12.764 | -34.5745 | -20.9165
12.7701 | -34.5749 | -20.9248
12.7762 | -34.5754 | -20.9331
12.7823 | -34.5759 | -20.9415
12.7884 | -34.5764 | -20.9499
12.7945 | -34.5769 | -20.9585
12.8006 | -34.5774 | -20.9668
12.8067 | -34.5778 | -20.9752
12.8128 | -34.5783 | -20.9834
12.8189 | -34.5788 | -20.9918
12.8251 | -34.5792 | -21.0003
12.8312 | -34.5797 | -21.0087
12.8373 | -34.5801 | -21.0169
12.8434 | -34.5806 | -21.0253
12.8495 | -34.5811 | -21.0335
12.8556 | -34.5815 | -21.0419
12.8617 | -34.582 | -21.0504
12.8678 | -34.5825 | -21.0589
12.8739 | -34.5829 | -21.0671
12.88 | -34.5834 | -21.0754
12.8861 | -34.5838 | -21.0836
12.8922 | -34.5843 | -21.092
12.8983 | -34.5848 | -21.1005
12.9044 | -34.5852 | -21.1087
12.9105 | -34.5857 | -21.117
12.9166 | -34.5862 | -21.1254
12.9227 | -34.5866 | -21.1337
12.9288 | -34.5871 | -21.1421
12.9349 | -34.5876 | -21.1506
12.941 | -34.5881 | -21.1589
12.9471 | -34.5885 | -21.1672
12.9532 | -34.589 | -21.1755
12.9593 | -34.5895 | -21.184
12.9654 | -34.59 | -21.1925
12.9716 | -34.5904 | -21.201
12.9777 | -34.5909 | -21.2094
12.9838 | -34.5914 | -21.2177
12.9899 | -34.5919 | -21.226
12.996 | -34.5923 | -21.2343
13.0021 | -34.5928 | -21.2427
13.0082 | -34.5933 | -21.2509
13.0143 | -34.5938 | -21.259
13.0204 | -34.5942 | -21.2673
13.0265 | -34.5947 | -21.2756
13.0326 | -34.5952 | -21.284
13.0387 | -34.5957 | -21.2924
13.0448 | -34.5961 | -21.3007
13.0509 | -34.5966 | -21.309
13.057 | -34.5971 | -21.3172
13.0631 | -34.5976 | -21.3254
13.0692 | -34.5981 | -21.3336
13.0753 | -34.5985 | -21.3421
13.0814 | -34.599 | -21.3506
13.0875 | -34.5995 | -21.3588
13.0936 | -34.5999 | -21.3672
13.0997 | -34.6004 | -21.3754
13.1058 | -34.6009 | -21.3837
13.112 | -34.6013 | -21.392
13.1181 | -34.6018 | -21.4002
13.1242 | -34.6023 | -21.4084
13.1303 | -34.6028 | -21.4168
13.1364 | -34.6032 | -21.4251
13.1425 | -34.6037 | -21.4333
13.1486 | -34.6041 | -21.4416
13.1547 | -34.6046 | -21.45
13.1608 | -34.6051 | -21.4583
13.1669 | -34.6056 | -21.4665
13.173 | -34.606 | -21.4748
13.1791 | -34.6065 | -21.4831
13.1852 | -34.607 | -21.4912
13.1913 | -34.6074 | -21.4993
13.1974 | -34.6079 | -21.5077
13.2035 | -34.6083 | -21.5161
13.2096 | -34.6088 | -21.5244
13.2157 | -34.6093 | -21.5326
13.2218 | -34.6098 | -21.5407
13.2279 | -34.6102 | -21.5488
13.234 | -34.6107 | -21.557
13.2401 | -34.6112 | -21.5652
13.2462 | -34.6116 | -21.5735
13.2524 | -34.6121 | -21.5818
13.2585 | -34.6125 | -21.5901
13.2646 | -34.613 | -21.5983
13.2707 | -34.6135 | -21.6066
13.2768 | -34.614 | -21.615
13.2829 | -34.6144 | -21.6234
13.289 | -34.6149 | -21.6314
13.2951 | -34.6154 | -21.6393
13.3012 | -34.6158 | -21.6475
13.3073 | -34.6163 | -21.6559
13.3134 | -34.6167 | -21.6642
13.3195 | -34.6172 | -21.6723
13.3256 | -34.6177 | -21.6805
13.3317 | -34.6181 | -21.6889
13.3378 | -34.6186 | -21.6973
13.3439 | -34.6191 | -21.7056
13.35 | -34.6195 | -21.7139
13.3561 | -34.62 | -21.722
13.3622 | -34.6204 | -21.7301
13.3683 | -34.6209 | -21.7383
13.3744 | -34.6214 | -21.7464
13.3805 | -34.6218 | -21.7547
13.3866 | -34.6223 | -21.7628
13.3927 | -34.6227 | -21.7711
13.3989 | -34.6232 | -21.7794
13.405 | -34.6236 | -21.7876
13.4111 | -34.6241 | -21.7958
13.4172 | -34.6245 | -21.8041
13.4233 | -34.625 | -21.8126
13.4294 | -34.6255 | -21.821
13.4355 | -34.6259 | -21.8291
13.4416 | -34.6264 | -21.8373
13.4477 | -34.6269 | -21.8455
13.4538 | -34.6273 | -21.8538
13.4599 | -34.6278 | -21.8621
13.466 | -34.6283 | -21.8705
13.4721 | -34.6287 | -21.8787
13.4782 | -34.6292 | -21.8871
13.4843 | -34.6296 | -21.8954
13.4904 | -34.6301 | -21.9035
13.4965 | -34.6305 | -21.9119
13.5026 | -34.631 | -21.9202
13.5087 | -34.6314 | -21.9282
13.5148 | -34.6319 | -21.9366
13.5209 | -34.6323 | -21.945
13.527 | -34.6328 | -21.9532
13.5331 | -34.6333 | -21.9614
13.5393 | -34.6338 | -21.9697
13.5454 | -34.6342 | -21.9779
13.5515 | -34.6347 | -21.9863
13.5576 | -34.6351 | -21.9946
13.5637 | -34.6356 | -22.003
13.5698 | -34.636 | -22.0112
13.5759 | -34.6365 | -22.0191
13.582 | -34.637 | -22.0273
13.5881 | -34.6374 | -22.0355
13.5942 | -34.6379 | -22.0439
13.6003 | -34.6384 | -22.052
13.6064 | -34.6388 | -22.0604
13.6125 | -34.6393 | -22.0688
13.6186 | -34.6397 | -22.077
13.6247 | -34.6402 | -22.0853
13.6308 | -34.6407 | -22.0937
13.6369 | -34.6411 | -22.1019
13.643 | -34.6415 | -22.1101
13.6491 | -34.642 | -22.1183
13.6552 | -34.6425 | -22.1265
13.6613 | -34.6429 | -22.1348
13.6674 | -34.6434 | -22.143
13.6735 | -34.6438 | -22.1512
13.6796 | -34.6443 | -22.1595
13.6858 | -34.6447 | -22.1678
13.6919 | -34.6452 | -22.1762
13.698 | -34.6456 | -22.1845
13.7041 | -34.6461 | -22.1929
13.7102 | -34.6466 | -22.2012
13.7163 | -34.647 | -22.2095
13.7224 | -34.6475 | -22.2177
13.7285 | -34.6479 | -22.226
13.7346 | -34.6484 | -22.2343
13.7407 | -34.6489 | -22.2426
13.7468 | -34.6493 | -22.2508
13.7529 | -34.6498 | -22.259
13.759 | -34.6502 | -22.2672
13.7651 | -34.6507 | -22.2754
13.7712 | -34.6511 | -22.2837
13.7773 | -34.6516 | -22.292
13.7834 | -34.652 | -22.3002
13.7895 | -34.6525 | -22.3085
13.7956 | -34.6529 | -22.3168
13.8017 | -34.6533 | -22.3249
13.8078 | -34.6538 | -22.3331
13.8139 | -34.6543 | -22.3412
13.82 | -34.6547 | -22.3492
13.8262 | -34.6551 | -22.3576
13.8323 | -34.6556 | -22.3659
13.8384 | -34.6561 | -22.3742
13.8445 | -34.6565 | -22.3825
13.8506 | -34.657 | -22.3908
13.8567 | -34.6574 | -22.3991
13.8628 | -34.6579 | -22.4075
13.8689 | -34.6584 | -22.4158
13.875 | -34.6588 | -22.4241
13.8811 | -34.6592 | -22.4323
13.8872 | -34.6596 | -22.4404
13.8933 | -34.6601 | -22.4486
13.8994 | -34.6605 | -22.4569
13.9055 | -34.661 | -22.4653
13.9116 | -34.6614 | -22.4735
13.9177 | -34.6619 | -22.4818
13.9238 | -34.6623 | -22.49
13.9299 | -34.6628 | -22.4983
13.936 | -34.6632 | -22.5066
13.9421 | -34.6636 | -22.5148
13.9482 | -34.6641 | -22.5229
13.9543 | -34.6645 | -22.5312
13.9604 | -34.665 | -22.5395
13.9665 | -34.6654 | -22.5476
13.9727 | -34.6659 | -22.5561
13.9788 | -34.6664 | -22.5644
13.9849 | -34.6668 | -22.5727
13.991 | -34.6673 | -22.581
13.9971 | -34.6677 | -22.589
14.0032 | -34.6681 | -22.5973
14.0093 | -34.6686 | -22.6055
14.0154 | -34.669 | -22.6137
14.0215 | -34.6695 | -22.6218
14.0276 | -34.67 | -22.6301
14.0337 | -34.6704 | -22.6385
14.0398 | -34.6709 | -22.6466
14.0459 | -34.6713 | -22.6548
14.052 | -34.6718 | -22.6632
14.0581 | -34.6722 | -22.6715
14.0642 | -34.6727 | -22.6797
14.0703 | -34.6731 | -22.688
14.0764 | -34.6735 | -22.6961
14.0825 | -34.674 | -22.7044
14.0886 | -34.6745 | -22.7127
14.0947 | -34.6749 | -22.721
14.1008 | -34.6754 | -22.7292
14.1069 | -34.6758 | -22.7373
14.1131 | -34.6762 | -22.7456
14.1192 | -34.6767 | -22.7538
14.1253 | -34.6771 | -22.7619
14.1314 | -34.6776 | -22.7702
14.1375 | -34.6781 | -22.7784
14.1436 | -34.6785 | -22.7866
14.1497 | -34.679 | -22.7948
14.1558 | -34.6794 | -22.803
14.1619 | -34.6799 | -22.8112
14.168 | -34.6803 | -22.8194
14.1741 | -34.6808 | -22.8277
14.1802 | -34.6812 | -22.8358
14.1863 | -34.6817 | -22.8441
14.1924 | -34.6821 | -22.8523
14.1985 | -34.6826 | -22.8605
14.2046 | -34.6831 | -22.8688
14.2107 | -34.6835 | -22.8768
14.2168 | -34.684 | -22.885
14.2229 | -34.6844 | -22.8932
14.229 | -34.6849 | -22.9016
14.2351 | -34.6853 | -22.9097
14.2412 | -34.6857 | -22.918
14.2473 | -34.6862 | -22.9263
14.2534 | -34.6866 | -22.9343
14.2596 | -34.6871 | -22.9423
14.2657 | -34.6875 | -22.9507
14.2718 | -34.6879 | -22.959
14.2779 | -34.6884 | -22.9673
14.284 | -34.6889 | -22.9755
14.2901 | -34.6893 | -22.9837
14.2962 | -34.6898 | -22.9919
14.3023 | -34.6902 | -23.0001
14.3084 | -34.6906 | -23.0084
14.3145 | -34.6911 | -23.0166
14.3206 | -34.6915 | -23.0251
14.3267 | -34.692 | -23.0332
14.3328 | -34.6924 | -23.0415
14.3389 | -34.6929 | -23.0497
14.345 | -34.6933 | -23.0579
14.3511 | -34.6937 | -23.0663
14.3572 | -34.6942 | -23.0746
14.3633 | -34.6946 | -23.0829
14.3694 | -34.6951 | -23.0911
14.3755 | -34.6956 | -23.0993
14.3816 | -34.696 | -23.1077
14.3877 | -34.6965 | -23.1159
14.3938 | -34.6969 | -23.1244
14.4 | -34.6974 | -23.1326
14.4061 | -34.6978 | -23.1407
14.4122 | -34.6983 | -23.1491
14.4183 | -34.6987 | -23.1571
14.4244 | -34.6992 | -23.1654
14.4305 | -34.6996 | -23.1737
14.4366 | -34.7001 | -23.1819
14.4427 | -34.7005 | -23.1903
14.4488 | -34.701 | -23.1986
14.4549 | -34.7015 | -23.2068
14.461 | -34.7019 | -23.2149
14.4671 | -34.7024 | -23.2233
14.4732 | -34.7028 | -23.2316
14.4793 | -34.7033 | -23.2399
14.4854 | -34.7037 | -23.248
14.4915 | -34.7042 | -23.2561
14.4976 | -34.7046 | -23.2642
14.5037 | -34.7051 | -23.2724
14.5098 | -34.7055 | -23.2807
14.5159 | -34.706 | -23.2889
14.522 | -34.7064 | -23.2971
14.5281 | -34.7069 | -23.3052
14.5342 | -34.7073 | -23.3135
14.5403 | -34.7078 | -23.3217
14.5465 | -34.7082 | -23.3299
14.5526 | -34.7087 | -23.3383
14.5587 | -34.7091 | -23.3466
14.5648 | -34.7096 | -23.3547
14.5709 | -34.7101 | -23.3629
14.577 | -34.7105 | -23.3711
14.5831 | -34.7109 | -23.3793
14.5892 | -34.7114 | -23.3875
14.5953 | -34.7118 | -23.3956
14.6014 | -34.7123 | -23.4039
14.6075 | -34.7127 | -23.412
14.6136 | -34.7131 | -23.4203
14.6197 | -34.7136 | -23.4286
14.6258 | -34.714 | -23.4368
14.6319 | -34.7145 | -23.445
14.638 | -34.7149 | -23.4532
14.6441 | -34.7154 | -23.4616
14.6502 | -34.7158 | -23.47
14.6563 | -34.7163 | -23.4783
14.6624 | -34.7167 | -23.4865
14.6685 | -34.7172 | -23.4945
14.6746 | -34.7176 | -23.5027
14.6807 | -34.7181 | -23.5109
14.6869 | -34.7186 | -23.5191
14.693 | -34.719 | -23.5272
14.6991 | -34.7195 | -23.5355
14.7052 | -34.72 | -23.5435
14.7113 | -34.7204 | -23.5518
14.7174 | -34.7209 | -23.5601
14.7235 | -34.7213 | -23.5682
14.7296 | -34.7218 | -23.5764
14.7357 | -34.7223 | -23.5845
14.7418 | -34.7227 | -23.5927
14.7479 | -34.7232 | -23.6008
14.754 | -34.7236 | -23.6091
14.7601 | -34.7241 | -23.6172
14.7662 | -34.7245 | -23.6254
14.7723 | -34.725 | -23.6334
14.7784 | -34.7254 | -23.6417
14.7845 | -34.7259 | -23.6499
14.7906 | -34.7263 | -23.6581
14.7967 | -34.7268 | -23.6662
14.8028 | -34.7272 | -23.6744
14.8089 | -34.7277 | -23.6825
14.815 | -34.7282 | -23.6907
14.8211 | -34.7286 | -23.6988
14.8272 | -34.7291 | -23.7071
14.8334 | -34.7295 | -23.7152
14.8395 | -34.73 | -23.7233
14.8456 | -34.7305 | -23.7315
14.8517 | -34.7309 | -23.7398
14.8578 | -34.7314 | -23.748
14.8639 | -34.7319 | -23.7563
14.87 | -34.7323 | -23.7645
14.8761 | -34.7328 | -23.7726
14.8822 | -34.7332 | -23.7809
14.8883 | -34.7337 | -23.7891
14.8944 | -34.7341 | -23.7971
14.9005 | -34.7346 | -23.8053
14.9066 | -34.735 | -23.8135
14.9127 | -34.7355 | -23.8215
14.9188 | -34.736 | -23.8297
14.9249 | -34.7364 | -23.8378
14.931 | -34.7369 | -23.846
14.9371 | -34.7374 | -23.8541
14.9432 | -34.7378 | -23.8623
14.9493 | -34.7383 | -23.8703
14.9554 | -34.7387 | -23.8784
14.9615 | -34.7392 | -23.8866
14.9676 | -34.7396 | -23.895
14.9738 | -34.7401 | -23.903
14.9799 | -34.7405 | -23.9113
14.986 | -34.741 | -23.9193
14.9921 | -34.7414 | -23.9275
14.9982 | -34.7418 | -23.9356
15.0043 | -34.7423 | -23.9436
15.0104 | -34.7428 | -23.9519
15.0165 | -34.7432 | -23.9601
15.0226 | -34.7437 | -23.9681
15.0287 | -34.7442 | -23.9763
15.0348 | -34.7446 | -23.9845
15.0409 | -34.7451 | -23.9927
15.047 | -34.7455 | -24.0009
15.0531 | -34.746 | -24.0089
15.0592 | -34.7465 | -24.017
15.0653 | -34.747 | -24.0252
15.0714 | -34.7474 | -24.0332
15.0775 | -34.7479 | -24.0413
15.0836 | -34.7483 | -24.0494
15.0897 | -34.7488 | -24.0573
15.0958 | -34.7493 | -24.0654
15.1019 | -34.7498 | -24.0736
15.108 | -34.7502 | -24.0816
15.1141 | -34.7507 | -24.0898
15.1203 | -34.7512 | -24.098
15.1264 | -34.7516 | -24.1061
15.1325 | -34.7521 | -24.1142
15.1386 | -34.7526 | -24.1223
15.1447 | -34.753 | -24.1304
15.1508 | -34.7535 | -24.1386
15.1569 | -34.754 | -24.1468
15.163 | -34.7545 | -24.1547
15.1691 | -34.7549 | -24.1626
15.1752 | -34.7554 | -24.1708
15.1813 | -34.7559 | -24.1789
15.1874 | -34.7563 | -24.187
15.1935 | -34.7568 | -24.1949
15.1996 | -34.7573 | -24.2029
15.2057 | -34.7577 | -24.211
15.2118 | -34.7582 | -24.2191
15.2179 | -34.7587 | -24.2274
15.224 | -34.7591 | -24.2355
15.2301 | -34.7596 | -24.2433
15.2362 | -34.7601 | -24.2515
15.2423 | -34.7606 | -24.2597
15.2484 | -34.761 | -24.2679
15.2545 | -34.7615 | -24.276
15.2607 | -34.762 | -24.2841
15.2668 | -34.7625 | -24.2919
15.2729 | -34.7629 | -24.3
15.279 | -34.7634 | -24.3083
15.2851 | -34.7639 | -24.3163
15.2912 | -34.7643 | -24.3244
15.2973 | -34.7648 | -24.3325
15.3034 | -34.7653 | -24.3406
15.3095 | -34.7658 | -24.3487
15.3156 | -34.7662 | -24.3567
15.3217 | -34.7667 | -24.3646
15.3278 | -34.7672 | -24.3728
15.3339 | -34.7677 | -24.3807
15.34 | -34.7682 | -24.3886
15.3461 | -34.7687 | -24.3966
15.3522 | -34.7691 | -24.4046
15.3583 | -34.7696 | -24.4128
15.3644 | -34.7701 | -24.4207
15.3705 | -34.7706 | -24.4285
15.3766 | -34.771 | -24.4366
15.3827 | -34.7715 | -24.4446
15.3888 | -34.772 | -24.4527
15.3949 | -34.7724 | -24.4607
15.401 | -34.7729 | -24.4688
15.4072 | -34.7734 | -24.4766
15.4133 | -34.7738 | -24.4847
15.4194 | -34.7743 | -24.4927
15.4255 | -34.7747 | -24.5007
15.4316 | -34.7752 | -24.5087
15.4377 | -34.7757 | -24.5169
15.4438 | -34.7762 | -24.5247
15.4499 | -34.7766 | -24.5326
15.456 | -34.7771 | -24.5406
15.4621 | -34.7776 | -24.5486
15.4682 | -34.7781 | -24.5567
15.4743 | -34.7785 | -24.5645
15.4804 | -34.779 | -24.5726
15.4865 | -34.7795 | -24.5806
15.4926 | -34.7799 | -24.5886
15.4987 | -34.7804 | -24.5967
15.5048 | -34.7809 | -24.6047
15.5109 | -34.7814 | -24.6126
15.517 | -34.7819 | -24.6207
15.5231 | -34.7824 | -24.6287
15.5292 | -34.7828 | -24.6366
15.5353 | -34.7833 | -24.6447
15.5414 | -34.7838 | -24.6526
15.5476 | -34.7843 | -24.6603
15.5537 | -34.7848 | -24.6683
15.5598 | -34.7852 | -24.6764
15.5659 | -34.7857 | -24.6843
15.572 | -34.7862 | -24.6922
15.5781 | -34.7867 | -24.7
15.5842 | -34.7872 | -24.7078
15.5903 | -34.7877 | -24.7156
15.5964 | -34.7882 | -24.7235
15.6025 | -34.7886 | -24.7314
15.6086 | -34.7891 | -24.7393
15.6147 | -34.7896 | -24.7473
15.6208 | -34.7901 | -24.7551
15.6269 | -34.7906 | -24.763
15.633 | -34.791 | -24.771
15.6391 | -34.7915 | -24.779
15.6452 | -34.792 | -24.787
15.6513 | -34.7925 | -24.7946
15.6574 | -34.793 | -24.8026
15.6635 | -34.7935 | -24.8105
15.6696 | -34.794 | -24.8186
15.6757 | -34.7944 | -24.8266
15.6818 | -34.7949 | -24.8345
15.688 | -34.7954 | -24.8423
15.6941 | -34.7958 | -24.8502
15.7002 | -34.7963 | -24.858
15.7063 | -34.7968 | -24.866
15.7124 | -34.7972 | -24.8741
15.7185 | -34.7977 | -24.882
15.7246 | -34.7982 | -24.8897
15.7307 | -34.7987 | -24.8975
15.7368 | -34.7992 | -24.9054
15.7429 | -34.7997 | -24.9131
15.749 | -34.8002 | -24.921
15.7551 | -34.8007 | -24.929
15.7612 | -34.8012 | -24.9368
15.7673 | -34.8017 | -24.9445
15.7734 | -34.8022 | -24.9523
15.7795 | -34.8026 | -24.9602
15.7856 | -34.8031 | -24.968
15.7917 | -34.8036 | -24.9756
15.7978 | -34.8041 | -24.9836
15.8039 | -34.8046 | -24.9916
15.81 | -34.8051 | -24.9994
15.8161 | -34.8055 | -25.0072
15.8222 | -34.806 | -25.015
15.8283 | -34.8065 | -25.0226
15.8345 | -34.807 | -25.0304
15.8406 | -34.8075 | -25.0383
15.8467 | -34.808 | -25.0463
15.8528 | -34.8085 | -25.0541
15.8589 | -34.809 | -25.062
15.865 | -34.8095 | -25.0694
15.8711 | -34.81 | -25.0773
15.8772 | -34.8105 | -25.085
15.8833 | -34.811 | -25.093
15.8894 | -34.8114 | -25.1011
15.8955 | -34.8119 | -25.1089
15.9016 | -34.8124 | -25.1164
15.9077 | -34.8129 | -25.1243
15.9138 | -34.8133 | -25.1321
15.9199 | -34.8138 | -25.1398
15.926 | -34.8143 | -25.1475
15.9321 | -34.8148 | -25.1553
15.9382 | -34.8153 | -25.1629
15.9443 | -34.8158 | -25.1708
15.9504 | -34.8163 | -25.1787
15.9565 | -34.8167 | -25.1865
15.9626 | -34.8172 | -25.1945
15.9687 | -34.8177 | -25.202
15.9749 | -34.8181 | -25.2096
15.981 | -34.8186 | -25.2175
15.9871 | -34.8191 | -25.2253
15.9932 | -34.8196 | -25.233
15.9993 | -34.8201 | -25.2409
16.0054 | -34.8206 | -25.2484
16.0115 | -34.8211 | -25.2563
16.0176 | -34.8216 | -25.2641
16.0237 | -34.8221 | -25.2719
16.0298 | -34.8225 | -25.2795
16.0359 | -34.823 | -25.2873
16.042 | -34.8235 | -25.2948
16.0481 | -34.824 | -25.3025
16.0542 | -34.8245 | -25.3103
16.0603 | -34.825 | -25.318
16.0664 | -34.8254 | -25.3259
16.0725 | -34.8259 | -25.3335
16.0786 | -34.8264 | -25.3411
16.0847 | -34.8269 | -25.349
16.0908 | -34.8274 | -25.3567
16.0969 | -34.8279 | -25.3645
16.103 | -34.8284 | -25.3723
16.1091 | -34.8289 | -25.3798
16.1152 | -34.8294 | -25.3876
16.1214 | -34.8299 | -25.3953
16.1275 | -34.8303 | -25.403
16.1336 | -34.8308 | -25.4107
16.1397 | -34.8314 | -25.4184
16.1458 | -34.8319 | -25.4257
16.1519 | -34.8324 | -25.4336
16.158 | -34.8328 | -25.4415
16.1641 | -34.8333 | -25.4494
16.1702 | -34.8338 | -25.4571
16.1763 | -34.8343 | -25.4648
16.1824 | -34.8349 | -25.4723
16.1885 | -34.8353 | -25.48
16.1946 | -34.8359 | -25.4877
16.2007 | -34.8363 | -25.4954
16.2068 | -34.8369 | -25.5032
16.2129 | -34.8373 | -25.511
16.219 | -34.8379 | -25.5184
16.2251 | -34.8384 | -25.5259
16.2312 | -34.8389 | -25.5335
16.2373 | -34.8394 | -25.5409
16.2434 | -34.8398 | -25.5487
16.2495 | -34.8403 | -25.5561
16.2556 | -34.8408 | -25.5639
16.2618 | -34.8413 | -25.5716
16.2679 | -34.8418 | -25.5794
16.274 | -34.8423 | -25.5872
16.2801 | -34.8428 | -25.5948
16.2862 | -34.8433 | -25.6025
16.2923 | -34.8438 | -25.6102
16.2984 | -34.8442 | -25.6179
16.3045 | -34.8447 | -25.6256
16.3106 | -34.8452 | -25.6334
16.3167 | -34.8457 | -25.6411
16.3228 | -34.8462 | -25.6485
16.3289 | -34.8467 | -25.656
16.335 | -34.8472 | -25.6637
16.3411 | -34.8477 | -25.6714
16.3472 | -34.8482 | -25.6792
16.3533 | -34.8487 | -25.6868
16.3594 | -34.8492 | -25.6942
16.3655 | -34.8497 | -25.7018
16.3716 | -34.8501 | -25.7094
16.3777 | -34.8507 | -25.7171
16.3838 | -34.8511 | -25.7249
16.3899 | -34.8516 | -25.7324
16.396 | -34.8521 | -25.74
16.4021 | -34.8526 | -25.7474
16.4083 | -34.8531 | -25.7551
16.4144 | -34.8536 | -25.7628
16.4205 | -34.8541 | -25.7704
16.4266 | -34.8546 | -25.7778
16.4327 | -34.855 | -25.7855
16.4388 | -34.8555 | -25.7931
16.4449 | -34.856 | -25.8007
16.451 | -34.8565 | -25.8083
16.4571 | -34.857 | -25.8161
16.4632 | -34.8575 | -25.8233
16.4693 | -34.8579 | -25.8309
16.4754 | -34.8584 | -25.8386
16.4815 | -34.8589 | -25.8463
16.4876 | -34.8594 | -25.8538
16.4937 | -34.8599 | -25.8613
16.4998 | -34.8604 | -25.8687
16.5059 | -34.8609 | -25.8761
16.512 | -34.8614 | -25.8838
16.5181 | -34.8619 | -25.8917
16.5242 | -34.8624 | -25.8993
16.5303 | -34.8629 | -25.9067
16.5364 | -34.8634 | -25.9142
16.5425 | -34.8639 | -25.9219
16.5487 | -34.8644 | -25.9293
16.5548 | -34.8649 | -25.9371
16.5609 | -34.8654 | -25.9446
16.567 | -34.8659 | -25.9519
16.5731 | -34.8664 | -25.9595
16.5792 | -34.8669 | -25.9672
16.5853 | -34.8674 | -25.9748
16.5914 | -34.8679 | -25.9825
16.5975 | -34.8684 | -25.9901
16.6036 | -34.8688 | -25.9977
16.6097 | -34.8693 | -26.0053
16.6158 | -34.8698 | -26.0129
16.6219 | -34.8703 | -26.0204
16.628 | -34.8708 | -26.028
16.6341 | -34.8713 | -26.0355
16.6402 | -34.8718 | -26.043
16.6463 | -34.8722 | -26.0505
16.6524 | -34.8727 | -26.058
16.6585 | -34.8732 | -26.0655
16.6646 | -34.8738 | -26.073
16.6707 | -34.8742 | -26.0805
16.6768 | -34.8747 | -26.0882
16.6829 | -34.8752 | -26.0955
16.689 | -34.8757 | -26.1031
16.6952 | -34.8762 | -26.1106
16.7013 | -34.8767 | -26.118
16.7074 | -34.8772 | -26.1255
16.7135 | -34.8777 | -26.1331
16.7196 | -34.8782 | -26.1406
16.7257 | -34.8787 | -26.148
16.7318 | -34.8792 | -26.1555
16.7379 | -34.8797 | -26.1628
16.744 | -34.8802 | -26.1701
16.7501 | -34.8807 | -26.1776
16.7562 | -34.8811 | -26.1851
16.7623 | -34.8816 | -26.1925
16.7684 | -34.8821 | -26.2002
16.7745 | -34.8826 | -26.2078
16.7806 | -34.8831 | -26.2149
16.7867 | -34.8837 | -26.2225
16.7928 | -34.8841 | -26.2299
16.7989 | -34.8846 | -26.2373
16.805 | -34.8851 | -26.2448
16.8111 | -34.8856 | -26.2521
16.8172 | -34.8861 | -26.2598
16.8233 | -34.8866 | -26.2674
16.8294 | -34.8871 | -26.2748
16.8356 | -34.8876 | -26.2825
16.8417 | -34.8881 | -26.29
16.8478 | -34.8886 | -26.2974
16.8539 | -34.8891 | -26.3047
16.86 | -34.8896 | -26.3124
16.8661 | -34.8901 | -26.3198
16.8722 | -34.8906 | -26.3275
16.8783 | -34.8911 | -26.335
16.8844 | -34.8916 | -26.3424
16.8905 | -34.8921 | -26.35
16.8966 | -34.8925 | -26.3574
16.9027 | -34.8931 | -26.365
16.9088 | -34.8936 | -26.3726
16.9149 | -34.8941 | -26.3802
16.921 | -34.8946 | -26.3874
16.9271 | -34.895 | -26.395
16.9332 | -34.8955 | -26.4026
16.9393 | -34.896 | -26.4102
16.9454 | -34.8965 | -26.4175
16.9515 | -34.8969 | -26.4251
16.9576 | -34.8975 | -26.4323
16.9637 | -34.8979 | -26.4395
16.9698 | -34.8985 | -26.4469
16.9759 | -34.8989 | -26.4543
16.9821 | -34.8994 | -26.462
16.9882 | -34.8999 | -26.4693
16.9943 | -34.9004 | -26.4767
17.0004 | -34.9009 | -26.4841
17.0065 | -34.9014 | -26.4916
17.0126 | -34.9019 | -26.4991
17.0187 | -34.9024 | -26.5067
17.0248 | -34.9029 | -26.514
17.0309 | -34.9034 | -26.5214
17.037 | -34.9039 | -26.5289
17.0431 | -34.9044 | -26.5362
17.0492 | -34.9048 | -26.5436
17.0553 | -34.9053 | -26.551
17.0614 | -34.9058 | -26.5585
17.0675 | -34.9063 | -26.5659
17.0736 | -34.9068 | -26.5734
17.0797 | -34.9073 | -26.5808
17.0858 | -34.9078 | -26.5885
17.0919 | -34.9083 | -26.596
17.098 | -34.9087 | -26.6032
17.1041 | -34.9092 | -26.6105
17.1102 | -34.9098 | -26.618
17.1163 | -34.9103 | -26.6254
17.1225 | -34.9108 | -26.6328
17.1286 | -34.9113 | -26.64
17.1347 | -34.9118 | -26.6475
17.1408 | -34.9123 | -26.6551
17.1469 | -34.9128 | -26.6626
17.153 | -34.9133 | -26.6699
17.1591 | -34.9137 | -26.6774
17.1652 | -34.9142 | -26.6849
17.1713 | -34.9147 | -26.6921
17.1774 | -34.9152 | -26.6996
17.1835 | -34.9157 | -26.7069
17.1896 | -34.9162 | -26.7144
17.1957 | -34.9167 | -26.7218
17.2018 | -34.9172 | -26.729
17.2079 | -34.9176 | -26.7366
17.214 | -34.9182 | -26.7442
17.2201 | -34.9186 | -26.7516
17.2262 | -34.9191 | -26.7592
17.2323 | -34.9196 | -26.7667
17.2384 | -34.9201 | -26.774
17.2445 | -34.9206 | -26.7813
17.2506 | -34.9211 | -26.7888
17.2567 | -34.9216 | -26.7963
17.2628 | -34.9221 | -26.8036
17.269 | -34.9226 | -26.811
17.2751 | -34.9231 | -26.8184
17.2812 | -34.9235 | -26.8259
17.2873 | -34.924 | -26.8333
17.2934 | -34.9245 | -26.8408
17.2995 | -34.925 | -26.8483
17.3056 | -34.9255 | -26.8554
17.3117 | -34.926 | -26.863
17.3178 | -34.9265 | -26.8702
17.3239 | -34.927 | -26.8776
17.33 | -34.9275 | -26.8848
17.3361 | -34.928 | -26.8922
17.3422 | -34.9285 | -26.8993
17.3483 | -34.929 | -26.9066
17.3544 | -34.9294 | -26.9141
17.3605 | -34.9299 | -26.9215
17.3666 | -34.9304 | -26.9289
17.3727 | -34.9309 | -26.9363
17.3788 | -34.9314 | -26.9434
17.3849 | -34.9319 | -26.9508
17.391 | -34.9324 | -26.9582
17.3971 | -34.9329 | -26.9656
17.4032 | -34.9334 | -26.9728
17.4094 | -34.9339 | -26.9802
17.4155 | -34.9343 | -26.9874
17.4216 | -34.9348 | -26.9947
17.4277 | -34.9353 | -27.002
17.4338 | -34.9358 | -27.0093
17.4399 | -34.9363 | -27.0168
17.446 | -34.9368 | -27.024
17.4521 | -34.9373 | -27.0314
17.4582 | -34.9378 | -27.0386
17.4643 | -34.9383 | -27.0458
17.4704 | -34.9388 | -27.0532
17.4765 | -34.9393 | -27.0606
17.4826 | -34.9398 | -27.0679
17.4887 | -34.9402 | -27.0752
17.4948 | -34.9407 | -27.0827
17.5009 | -34.9412 | -27.09
17.507 | -34.9417 | -27.0975
17.5131 | -34.9422 | -27.1048
17.5192 | -34.9427 | -27.112
17.5253 | -34.9432 | -27.1193
17.5314 | -34.9436 | -27.1267
17.5375 | -34.9441 | -27.1339
17.5436 | -34.9446 | -27.1414
17.5497 | -34.9451 | -27.1487
17.5559 | -34.9456 | -27.156
17.562 | -34.9461 | -27.1631
17.5681 | -34.9466 | -27.1705
17.5742 | -34.947 | -27.1779
17.5803 | -34.9475 | -27.1852
17.5864 | -34.948 | -27.1926
17.5925 | -34.9485 | -27.2001
17.5986 | -34.949 | -27.2074
17.6047 | -34.9495 | -27.2148
17.6108 | -34.95 | -27.2221
17.6169 | -34.9505 | -27.2293
17.623 | -34.9509 | -27.2363
17.6291 | -34.9514 | -27.2435
17.6352 | -34.9519 | -27.2507
17.6413 | -34.9524 | -27.2581
17.6474 | -34.9529 | -27.2654
17.6535 | -34.9534 | -27.2727
17.6596 | -34.9538 | -27.2799
17.6657 | -34.9544 | -27.2874
17.6718 | -34.9548 | -27.2946
17.6779 | -34.9553 | -27.3019
17.684 | -34.9558 | -27.3094
17.6901 | -34.9563 | -27.3166
17.6963 | -34.9567 | -27.3237
17.7024 | -34.9572 | -27.3309
17.7085 | -34.9577 | -27.338
17.7146 | -34.9582 | -27.3452
17.7207 | -34.9587 | -27.3524
17.7268 | -34.9592 | -27.3595
17.7329 | -34.9597 | -27.3669
17.739 | -34.9602 | -27.3742
17.7451 | -34.9607 | -27.3813
17.7512 | -34.9612 | -27.3886
17.7573 | -34.9616 | -27.396
17.7634 | -34.9621 | -27.403
17.7695 | -34.9626 | -27.4102
17.7756 | -34.9631 | -27.4175
17.7817 | -34.9636 | -27.4247
17.7878 | -34.9641 | -27.432
17.7939 | -34.9646 | -27.4393
17.8 | -34.9651 | -27.4466
17.8061 | -34.9655 | -27.4539
17.8122 | -34.966 | -27.4612
17.8183 | -34.9665 | -27.4685
17.8244 | -34.967 | -27.4756
17.8305 | -34.9674 | -27.483
17.8366 | -34.9679 | -27.4901
17.8428 | -34.9684 | -27.4975
17.8489 | -34.9689 | -27.5047
17.855 | -34.9694 | -27.5118
17.8611 | -34.9699 | -27.519
17.8672 | -34.9704 | -27.5262
17.8733 | -34.9709 | -27.5335
17.8794 | -34.9713 | -27.5408
17.8855 | -34.9718 | -27.5481
17.8916 | -34.9723 | -27.5555
17.8977 | -34.9727 | -27.5628
17.9038 | -34.9732 | -27.5699
17.9099 | -34.9737 | -27.5773
17.916 | -34.9742 | -27.5846
17.9221 | -34.9747 | -27.5918
17.9282 | -34.9752 | -27.599
17.9343 | -34.9757 | -27.6062
17.9404 | -34.9761 | -27.6132
17.9465 | -34.9766 | -27.6203
17.9526 | -34.9771 | -27.6276
17.9587 | -34.9775 | -27.6349
17.9648 | -34.978 | -27.642
17.9709 | -34.9785 | -27.6493
17.977 | -34.979 | -27.6566
17.9832 | -34.9794 | -27.6639
17.9893 | -34.9799 | -27.6713
17.9954 | -34.9804 | -27.6786
18.0015 | -34.9809 | -27.686
18.0076 | -34.9814 | -27.693
18.0137 | -34.9818 | -27.7003
18.0198 | -34.9823 | -27.7075
18.0259 | -34.9827 | -27.7148
18.032 | -34.9832 | -27.722
18.0381 | -34.9837 | -27.7292
18.0442 | -34.9842 | -27.7365
18.0503 | -34.9847 | -27.7438
18.0564 | -34.9851 | -27.7511
18.0625 | -34.9856 | -27.7585
18.0686 | -34.9861 | -27.7657
18.0747 | -34.9866 | -27.7729
18.0808 | -34.9871 | -27.78
18.0869 | -34.9875 | -27.7872
18.093 | -34.988 | -27.7945
18.0991 | -34.9885 | -27.8018
18.1052 | -34.9889 | -27.8089
18.1113 | -34.9894 | -27.8161
18.1174 | -34.9899 | -27.8232
18.1236 | -34.9904 | -27.8304
18.1297 | -34.9909 | -27.8377
18.1358 | -34.9914 | -27.8449
18.1419 | -34.9919 | -27.852
18.148 | -34.9923 | -27.8592
18.1541 | -34.9929 | -27.8661
18.1602 | -34.9933 | -27.8734
18.1663 | -34.9938 | -27.8806
18.1724 | -34.9943 | -27.8878
18.1785 | -34.9948 | -27.8949
18.1846 | -34.9952 | -27.902
18.1907 | -34.9957 | -27.9091
18.1968 | -34.9962 | -27.9164
18.2029 | -34.9967 | -27.9236
18.209 | -34.9971 | -27.9309
18.2151 | -34.9976 | -27.9382
18.2212 | -34.9981 | -27.9453
18.2273 | -34.9986 | -27.9524
18.2334 | -34.9991 | -27.9596
18.2395 | -34.9995 | -27.9666
18.2456 | -35 | -27.974
18.2517 | -35.0005 | -27.9812
18.2578 | -35.0009 | -27.988
18.2639 | -35.0014 | -27.9955
18.2701 | -35.0019 | -28.0025
18.2762 | -35.0024 | -28.0096
18.2823 | -35.0029 | -28.0167
18.2884 | -35.0034 | -28.0239
18.2945 | -35.0039 | -28.0309
18.3006 | -35.0043 | -28.0381
18.3067 | -35.0048 | -28.0454
18.3128 | -35.0053 | -28.0527
18.3189 | -35.0058 | -28.0598
18.325 | -35.0062 | -28.0666
18.3311 | -35.0067 | -28.0737
18.3372 | -35.0072 | -28.0808
18.3433 | -35.0077 | -28.088
18.3494 | -35.0081 | -28.0951
18.3555 | -35.0086 | -28.1023
18.3616 | -35.0091 | -28.1092
18.3677 | -35.0095 | -28.1164
18.3738 | -35.01 | -28.1235
18.3799 | -35.0105 | -28.1306
18.386 | -35.0109 | -28.1379
18.3921 | -35.0114 | -28.145
18.3982 | -35.0118 | -28.152
18.4043 | -35.0123 | -28.1592
18.4105 | -35.0128 | -28.1664
18.4166 | -35.0132 | -28.1736
18.4227 | -35.0137 | -28.1807
18.4288 | -35.0142 | -28.1878
18.4349 | -35.0147 | -28.1947
18.441 | -35.0151 | -28.2018
18.4471 | -35.0156 | -28.2089
18.4532 | -35.0161 | -28.2161
18.4593 | -35.0165 | -28.2231
18.4654 | -35.017 | -28.2299
18.4715 | -35.0174 | -28.237
18.4776 | -35.0179 | -28.2443
18.4837 | -35.0184 | -28.2514
18.4898 | -35.0188 | -28.2585
18.4959 | -35.0193 | -28.2658
18.502 | -35.0198 | -28.273
18.5081 | -35.0202 | -28.2801
18.5142 | -35.0207 | -28.2872
18.5203 | -35.0211 | -28.2943
18.5264 | -35.0216 | -28.3014
18.5325 | -35.0221 | -28.3086
18.5386 | -35.0225 | -28.3158
18.5447 | -35.023 | -28.3228
18.5508 | -35.0235 | -28.3299
18.557 | -35.024 | -28.3372
18.5631 | -35.0244 | -28.3443
18.5692 | -35.0249 | -28.3513
18.5753 | -35.0254 | -28.3582
18.5814 | -35.0258 | -28.3652
18.5875 | -35.0263 | -28.3723
18.5936 | -35.0268 | -28.3793
18.5997 | -35.0272 | -28.3866
18.6058 | -35.0277 | -28.3934
18.6119 | -35.0282 | -28.4004
18.618 | -35.0286 | -28.4076
18.6241 | -35.0291 | -28.4148
18.6302 | -35.0295 | -28.4222
18.6363 | -35.03 | -28.4294
18.6424 | -35.0304 | -28.4365
18.6485 | -35.0309 | -28.4435
18.6546 | -35.0313 | -28.4507
18.6607 | -35.0318 | -28.4578
18.6668 | -35.0322 | -28.4649
18.6729 | -35.0327 | -28.4721
18.679 | -35.0331 | -28.4789
18.6851 | -35.0336 | -28.4859
18.6912 | -35.0341 | -28.4929
18.6974 | -35.0345 | -28.4999
18.7035 | -35.035 | -28.5071
18.7096 | -35.0355 | -28.5143
18.7157 | -35.036 | -28.5214
18.7218 | -35.0364 | -28.5285
18.7279 | -35.0369 | -28.5356
18.734 | -35.0374 | -28.543
18.7401 | -35.0379 | -28.55
18.7462 | -35.0383 | -28.5569
18.7523 | -35.0388 | -28.564
18.7584 | -35.0393 | -28.5711
18.7645 | -35.0398 | -28.5782
18.7706 | -35.0402 | -28.5853
18.7767 | -35.0407 | -28.5923
18.7828 | -35.0411 | -28.5991
18.7889 | -35.0416 | -28.6063
18.795 | -35.0421 | -28.6135
18.8011 | -35.0425 | -28.6207
18.8072 | -35.043 | -28.628
18.8133 | -35.0435 | -28.6352
18.8194 | -35.0439 | -28.642
18.8255 | -35.0444 | -28.6489
18.8316 | -35.0448 | -28.6561
18.8377 | -35.0453 | -28.6631
18.8439 | -35.0458 | -28.6701
18.85 | -35.0462 | -28.6773
18.8561 | -35.0467 | -28.6842
18.8622 | -35.0471 | -28.6912
18.8683 | -35.0476 | -28.6981
18.8744 | -35.0481 | -28.705
18.8805 | -35.0485 | -28.7123
18.8866 | -35.049 | -28.7196
18.8927 | -35.0494 | -28.7265
18.8988 | -35.0499 | -28.7336
18.9049 | -35.0504 | -28.7406
18.911 | -35.0508 | -28.7478
18.9171 | -35.0513 | -28.755
18.9232 | -35.0518 | -28.7616
18.9293 | -35.0523 | -28.7686
18.9354 | -35.0527 | -28.7757
18.9415 | -35.0532 | -28.7826
18.9476 | -35.0537 | -28.7895
18.9537 | -35.0541 | -28.7966
18.9598 | -35.0545 | -28.8035
18.9659 | -35.055 | -28.8106
18.972 | -35.0555 | -28.8176
18.9781 | -35.0559 | -28.8246
18.9843 | -35.0564 | -28.8318
18.9904 | -35.0568 | -28.8388
18.9965 | -35.0573 | -28.8455
19.0026 | -35.0578 | -28.8526
19.0087 | -35.0582 | -28.8597
19.0148 | -35.0587 | -28.8668
19.0209 | -35.0592 | -28.8739
19.027 | -35.0597 | -28.881
19.0331 | -35.0601 | -28.8877
19.0392 | -35.0606 | -28.8946
19.0453 | -35.0611 | -28.9016
19.0514 | -35.0615 | -28.9086
19.0575 | -35.062 | -28.9156
19.0636 | -35.0624 | -28.9224
19.0697 | -35.0629 | -28.9294
19.0758 | -35.0633 | -28.9362
19.0819 | -35.0638 | -28.9432
19.088 | -35.0643 | -28.9503
19.0941 | -35.0647 | -28.9575
19.1002 | -35.0651 | -28.9644
19.1063 | -35.0656 | -28.9715
19.1124 | -35.066 | -28.9784
19.1185 | -35.0664 | -28.9855
19.1246 | -35.0669 | -28.9924
19.1308 | -35.0674 | -28.9994
19.1369 | -35.0678 | -29.0064
19.143 | -35.0683 | -29.0133
19.1491 | -35.0687 | -29.0205
19.1552 | -35.0692 | -29.0275
19.1613 | -35.0696 | -29.0345
19.1674 | -35.0701 | -29.0414
19.1735 | -35.0706 | -29.0484
19.1796 | -35.071 | -29.0555
19.1857 | -35.0714 | -29.0625
19.1918 | -35.0719 | -29.0695
19.1979 | -35.0723 | -29.0767
19.204 | -35.0728 | -29.0836
19.2101 | -35.0732 | -29.0906
19.2162 | -35.0737 | -29.0974
19.2223 | -35.0742 | -29.1043
19.2284 | -35.0746 | -29.1111
19.2345 | -35.075 | -29.1181
19.2406 | -35.0755 | -29.1249
19.2467 | -35.076 | -29.1318
19.2528 | -35.0764 | -29.1389
19.2589 | -35.0769 | -29.1459
19.265 | -35.0774 | -29.1529
19.2712 | -35.0778 | -29.1597
19.2773 | -35.0782 | -29.1662
19.2834 | -35.0787 | -29.1731
19.2895 | -35.0792 | -29.18
19.2956 | -35.0796 | -29.1871
19.3017 | -35.0801 | -29.1938
19.3078 | -35.0805 | -29.2009
19.3139 | -35.081 | -29.2076
19.32 | -35.0814 | -29.2146
19.3261 | -35.0819 | -29.2215
19.3322 | -35.0823 | -29.2284
19.3383 | -35.0828 | -29.2353
19.3444 | -35.0833 | -29.2419
19.3505 | -35.0837 | -29.2489
19.3566 | -35.0842 | -29.2557
19.3627 | -35.0847 | -29.2627
19.3688 | -35.0851 | -29.2695
19.3749 | -35.0856 | -29.2765
19.381 | -35.0861 | -29.2833
19.3871 | -35.0865 | -29.2902
19.3932 | -35.087 | -29.2971
19.3993 | -35.0874 | -29.304
19.4054 | -35.0878 | -29.3109
19.4115 | -35.0883 | -29.3179
19.4177 | -35.0888 | -29.3245
19.4238 | -35.0892 | -29.3315
19.4299 | -35.0897 | -29.3384
19.436 | -35.0902 | -29.3456
19.4421 | -35.0906 | -29.3523
19.4482 | -35.091 | -29.3592
19.4543 | -35.0915 | -29.3659
19.4604 | -35.092 | -29.3727
19.4665 | -35.0924 | -29.3796
19.4726 | -35.0929 | -29.3862
19.4787 | -35.0933 | -29.3931
19.4848 | -35.0938 | -29.4
19.4909 | -35.0943 | -29.4068
19.497 | -35.0947 | -29.4137
19.5031 | -35.0952 | -29.4208
19.5092 | -35.0956 | -29.4277
19.5153 | -35.096 | -29.4346
19.5214 | -35.0965 | -29.4414
19.5275 | -35.0969 | -29.4482
19.5336 | -35.0973 | -29.4552
19.5397 | -35.0978 | -29.462
19.5458 | -35.0982 | -29.469
19.5519 | -35.0987 | -29.4759
19.5581 | -35.0991 | -29.4829
19.5642 | -35.0995 | -29.4896
19.5703 | -35.1 | -29.4965
19.5764 | -35.1004 | -29.5032
19.5825 | -35.1008 | -29.5101
19.5886 | -35.1013 | -29.517
19.5947 | -35.1017 | -29.5235
19.6008 | -35.1022 | -29.5303
19.6069 | -35.1026 | -29.5373
19.613 | -35.1031 | -29.5446
19.6191 | -35.1036 | -29.5516
19.6252 | -35.104 | -29.5585
19.6313 | -35.1045 | -29.565
19.6374 | -35.1049 | -29.5721
19.6435 | -35.1053 | -29.5789
19.6496 | -35.1058 | -29.5856
19.6557 | -35.1063 | -29.5923
19.6618 | -35.1067 | -29.599
19.6679 | -35.1071 | -29.6059
19.674 | -35.1076 | -29.6126
19.6801 | -35.108 | -29.6194
19.6862 | -35.1085 | -29.6261
19.6923 | -35.109 | -29.633
19.6984 | -35.1094 | -29.6395
19.7046 | -35.1098 | -29.6463
19.7107 | -35.1103 | -29.6531
19.7168 | -35.1108 | -29.6599
19.7229 | -35.1112 | -29.667
19.729 | -35.1116 | -29.6741
19.7351 | -35.1121 | -29.6806
19.7412 | -35.1125 | -29.6877
19.7473 | -35.113 | -29.6944
19.7534 | -35.1134 | -29.7013
19.7595 | -35.1139 | -29.7082
19.7656 | -35.1143 | -29.715
19.7717 | -35.1147 | -29.7215
19.7778 | -35.1152 | -29.7282
19.7839 | -35.1156 | -29.735
19.79 | -35.1161 | -29.7416
19.7961 | -35.1165 | -29.7486
19.8022 | -35.1169 | -29.7551
19.8083 | -35.1174 | -29.762
19.8144 | -35.1178 | -29.769
19.8205 | -35.1183 | -29.7758
19.8266 | -35.1188 | -29.7824
19.8327 | -35.1192 | -29.7891
19.8388 | -35.1197 | -29.7956
19.845 | -35.1201 | -29.8023
19.8511 | -35.1206 | -29.8091
19.8572 | -35.121 | -29.8159
19.8633 | -35.1214 | -29.8229
19.8694 | -35.1219 | -29.8298
19.8755 | -35.1223 | -29.8363
19.8816 | -35.1228 | -29.8433
19.8877 | -35.1232 | -29.8503
19.8938 | -35.1236 | -29.8573
19.8999 | -35.1241 | -29.8642
19.906 | -35.1245 | -29.871
19.9121 | -35.125 | -29.8777
19.9182 | -35.1254 | -29.8844
19.9243 | -35.1258 | -29.8913
19.9304 | -35.1263 | -29.898
19.9365 | -35.1267 | -29.9049
19.9426 | -35.1272 | -29.9114
19.9487 | -35.1277 | -29.9182
19.9548 | -35.1281 | -29.925
19.9609 | -35.1285 | -29.9316
19.967 | -35.1289 | -29.9383
19.9731 | -35.1294 | -29.9451
19.9792 | -35.1298 | -29.9519
19.9853 | -35.1303 | -29.9586
19.9915 | -35.1307 | -29.9652
19.9976 | -35.1311 | -29.9719
20.0037 | -35.1316 | -29.9789
20.0098 | -35.132 | -29.9859
20.0159 | -35.1325 | -29.9926
20.022 | -35.1329 | -29.9992
20.0281 | -35.1334 | -30.006
20.0342 | -35.1338 | -30.0128
20.0403 | -35.1342 | -30.0197
20.0464 | -35.1347 | -30.0263
20.0525 | -35.1351 | -30.0327
20.0586 | -35.1355 | -30.0397
20.0647 | -35.136 | -30.0465
20.0708 | -35.1365 | -30.0534
20.0769 | -35.1369 | -30.06
20.083 | -35.1373 | -30.0664
20.0891 | -35.1377 | -30.0732
20.0952 | -35.1382 | -30.0799
20.1013 | -35.1387 | -30.0867
20.1074 | -35.1391 | -30.0938
20.1135 | -35.1395 | -30.1006
20.1196 | -35.1399 | -30.107
20.1257 | -35.1403 | -30.1138
20.1319 | -35.1408 | -30.1208
20.138 | -35.1412 | -30.1274
20.1441 | -35.1417 | -30.1339
20.1502 | -35.1421 | -30.1408
20.1563 | -35.1425 | -30.1476
20.1624 | -35.143 | -30.1542
20.1685 | -35.1434 | -30.1609
20.1746 | -35.1439 | -30.1678
20.1807 | -35.1443 | -30.1745
20.1868 | -35.1447 | -30.1811
20.1929 | -35.1452 | -30.1874
20.199 | -35.1456 | -30.1945
20.2051 | -35.146 | -30.2011
20.2112 | -35.1465 | -30.2079
20.2173 | -35.1469 | -30.2143
20.2234 | -35.1474 | -30.2213
20.2295 | -35.1478 | -30.228
20.2356 | -35.1483 | -30.2349
20.2417 | -35.1487 | -30.2418
20.2478 | -35.1491 | -30.2485
20.2539 | -35.1496 | -30.2551
20.26 | -35.15 | -30.2615
20.2661 | -35.1505 | -30.2683
20.2723 | -35.151 | -30.2751
20.2784 | -35.1514 | -30.2816
20.2845 | -35.1518 | -30.2883
20.2906 | -35.1522 | -30.295
20.2967 | -35.1527 | -30.3017
20.3028 | -35.1531 | -30.3085
20.3089 | -35.1536 | -30.3152
20.315 | -35.154 | -30.3219
20.3211 | -35.1545 | -30.3289
20.3272 | -35.1549 | -30.3357
20.3333 | -35.1554 | -30.3422
20.3394 | -35.1558 | -30.3487
20.3455 | -35.1562 | -30.3552
20.3516 | -35.1567 | -30.3617
20.3577 | -35.1571 | -30.3683
20.3638 | -35.1576 | -30.3751
20.3699 | -35.1581 | -30.3814
20.376 | -35.1585 | -30.3882
20.3821 | -35.1589 | -30.3951
20.3882 | -35.1594 | -30.4015
20.3943 | -35.1598 | -30.4083
20.4004 | -35.1602 | -30.4146
20.4065 | -35.1606 | -30.4212
20.4126 | -35.1611 | -30.4279
20.4188 | -35.1615 | -30.4348
20.4249 | -35.162 | -30.4412
20.431 | -35.1624 | -30.4479
20.4371 | -35.1629 | -30.4544
20.4432 | -35.1633 | -30.4608
20.4493 | -35.1638 | -30.4678
20.4554 | -35.1642 | -30.4743
20.4615 | -35.1647 | -30.481
20.4676 | -35.1651 | -30.4879
20.4737 | -35.1655 | -30.4944
20.4798 | -35.1659 | -30.5009
20.4859 | -35.1664 | -30.5081
20.492 | -35.1668 | -30.5148
20.4981 | -35.1673 | -30.5212
20.5042 | -35.1677 | -30.5278
20.5103 | -35.1681 | -30.5342
20.5164 | -35.1686 | -30.541
20.5225 | -35.169 | -30.5475
20.5286 | -35.1694 | -30.5543
20.5347 | -35.1699 | -30.561
20.5408 | -35.1703 | -30.5677
20.5469 | -35.1708 | -30.5746
20.553 | -35.1712 | -30.5813
20.5592 | -35.1717 | -30.5882
20.5653 | -35.1721 | -30.5948
20.5714 | -35.1725 | -30.6013
20.5775 | -35.173 | -30.6077
20.5836 | -35.1735 | -30.6143
20.5897 | -35.1739 | -30.621
20.5958 | -35.1744 | -30.6277
20.6019 | -35.1748 | -30.6343
20.608 | -35.1752 | -30.6408
20.6141 | -35.1757 | -30.6473
20.6202 | -35.1761 | -30.6542
20.6263 | -35.1765 | -30.6607
20.6324 | -35.177 | -30.6674
20.6385 | -35.1775 | -30.674
20.6446 | -35.1779 | -30.6808
20.6507 | -35.1784 | -30.6873
20.6568 | -35.1788 | -30.6938
20.6629 | -35.1793 | -30.7006
20.669 | -35.1797 | -30.7075
20.6751 | -35.1802 | -30.7141
20.6812 | -35.1806 | -30.7204
20.6873 | -35.1811 | -30.7272
20.6934 | -35.1815 | -30.7342
20.6995 | -35.1819 | -30.7408
20.7057 | -35.1824 | -30.7475
20.7118 | -35.1829 | -30.754
20.7179 | -35.1833 | -30.7604
20.724 | -35.1837 | -30.7674
20.7301 | -35.1842 | -30.7739
20.7362 | -35.1846 | -30.7806
20.7423 | -35.185 | -30.7872
20.7484 | -35.1855 | -30.7939
20.7545 | -35.186 | -30.8004
20.7606 | -35.1865 | -30.8069
20.7667 | -35.187 | -30.8134
20.7728 | -35.1874 | -30.82
20.7789 | -35.1878 | -30.8265
20.785 | -35.1883 | -30.833
20.7911 | -35.1887 | -30.8396
20.7972 | -35.1891 | -30.8464
20.8033 | -35.1896 | -30.8531
20.8094 | -35.19 | -30.8598
20.8155 | -35.1904 | -30.8665
20.8216 | -35.1909 | -30.8731
20.8277 | -35.1913 | -30.8796
20.8338 | -35.1917 | -30.8863
20.8399 | -35.1922 | -30.8931
20.8461 | -35.1926 | -30.8997
20.8522 | -35.1931 | -30.9062
20.8583 | -35.1935 | -30.9124
20.8644 | -35.194 | -30.9192
20.8705 | -35.1944 | -30.9263
20.8766 | -35.1948 | -30.9331
20.8827 | -35.1953 | -30.9396
20.8888 | -35.1957 | -30.946
20.8949 | -35.1962 | -30.9522
20.901 | -35.1966 | -30.9585
20.9071 | -35.1971 | -30.9648
20.9132 | -35.1976 | -30.9715
20.9193 | -35.198 | -30.9781
20.9254 | -35.1984 | -30.9846
20.9315 | -35.1988 | -30.991
20.9376 | -35.1992 | -30.9974
20.9437 | -35.1997 | -31.0038
20.9498 | -35.2001 | -31.0104
20.9559 | -35.2006 | -31.0168
20.962 | -35.2011 | -31.0234
20.9681 | -35.2015 | -31.03
20.9742 | -35.2019 | -31.0366
20.9803 | -35.2024 | -31.0429
20.9864 | -35.2028 | -31.0497
20.9926 | -35.2032 | -31.0563
20.9987 | -35.2037 | -31.0629
21.0048 | -35.2042 | -31.0694
21.0109 | -35.2046 | -31.0762
21.017 | -35.205 | -31.0826
21.0231 | -35.2055 | -31.0892
21.0292 | -35.206 | -31.0955
21.0353 | -35.2064 | -31.102
21.0414 | -35.2068 | -31.1086
21.0475 | -35.2072 | -31.1151
21.0536 | -35.2077 | -31.1213
21.0597 | -35.2082 | -31.1278
21.0658 | -35.2086 | -31.1343
21.0719 | -35.2091 | -31.141
21.078 | -35.2095 | -31.1477
21.0841 | -35.21 | -31.1547
21.0902 | -35.2105 | -31.161
21.0963 | -35.2109 | -31.1678
21.1024 | -35.2114 | -31.1742
21.1085 | -35.2118 | -31.1808
21.1146 | -35.2123 | -31.1877
21.1207 | -35.2127 | -31.1944
21.1268 | -35.2132 | -31.2012
21.133 | -35.2136 | -31.2078
21.1391 | -35.2141 | -31.2141
21.1452 | -35.2145 | -31.2208
21.1513 | -35.2149 | -31.2274
21.1574 | -35.2153 | -31.234
21.1635 | -35.2157 | -31.2404
21.1696 | -35.2162 | -31.2469
21.1757 | -35.2166 | -31.2532
21.1818 | -35.217 | -31.2599
21.1879 | -35.2175 | -31.2668
21.194 | -35.2179 | -31.2732
21.2001 | -35.2184 | -31.2796
21.2062 | -35.2188 | -31.286
21.2123 | -35.2193 | -31.2922
21.2184 | -35.2197 | -31.2987
21.2245 | -35.2202 | -31.3051
21.2306 | -35.2206 | -31.3118
21.2367 | -35.221 | -31.3185
21.2428 | -35.2214 | -31.3249
21.2489 | -35.2219 | -31.3315
21.255 | -35.2223 | -31.3381
21.2611 | -35.2227 | -31.3446
21.2672 | -35.2232 | -31.351
21.2733 | -35.2236 | -31.3576
21.2795 | -35.2241 | -31.364
21.2856 | -35.2246 | -31.3706
21.2917 | -35.2251 | -31.3771
21.2978 | -35.2255 | -31.3833
21.3039 | -35.226 | -31.3894
21.31 | -35.2264 | -31.3959
21.3161 | -35.2268 | -31.4024
21.3222 | -35.2272 | -31.4087
21.3283 | -35.2277 | -31.4151
21.3344 | -35.2281 | -31.4216
21.3405 | -35.2286 | -31.4281
21.3466 | -35.229 | -31.4345
21.3527 | -35.2295 | -31.441
21.3588 | -35.23 | -31.4474
21.3649 | -35.2304 | -31.4538
21.371 | -35.231 | -31.4601
21.3771 | -35.2314 | -31.4667
21.3832 | -35.2318 | -31.4731
21.3893 | -35.2322 | -31.4793
21.3954 | -35.2326 | -31.4859
21.4015 | -35.2331 | -31.4925
21.4076 | -35.2335 | -31.4987
21.4137 | -35.234 | -31.5056
21.4199 | -35.2344 | -31.5119
21.426 | -35.2349 | -31.5183
21.4321 | -35.2353 | -31.5248
21.4382 | -35.2358 | -31.5312
21.4443 | -35.2362 | -31.5374
21.4504 | -35.2366 | -31.544
21.4565 | -35.2371 | -31.55
21.4626 | -35.2376 | -31.5566
21.4687 | -35.2381 | -31.5631
21.4748 | -35.2385 | -31.5696
21.4809 | -35.239 | -31.5758
21.487 | -35.2394 | -31.5823
21.4931 | -35.2399 | -31.5885
21.4992 | -35.2404 | -31.5947
21.5053 | -35.2409 | -31.6012
21.5114 | -35.2413 | -31.6079
21.5175 | -35.2417 | -31.6144
21.5236 | -35.2421 | -31.6211
21.5297 | -35.2426 | -31.627
21.5358 | -35.2431 | -31.6333
21.5419 | -35.2436 | -31.6396
21.548 | -35.244 | -31.6462
21.5541 | -35.2444 | -31.6527
21.5602 | -35.2448 | -31.6591
21.5664 | -35.2453 | -31.6652
21.5725 | -35.2457 | -31.6719
21.5786 | -35.2461 | -31.6781
21.5847 | -35.2466 | -31.6846
21.5908 | -35.247 | -31.691
21.5969 | -35.2475 | -31.6975
21.603 | -35.248 | -31.7039
21.6091 | -35.2484 | -31.7103
21.6152 | -35.2488 | -31.7169
21.6213 | -35.2493 | -31.7236
21.6274 | -35.2497 | -31.7302
21.6335 | -35.2502 | -31.7365
21.6396 | -35.2507 | -31.7432
21.6457 | -35.251 | -31.7496
21.6518 | -35.2515 | -31.7563
21.6579 | -35.252 | -31.7626
21.664 | -35.2524 | -31.769
21.6701 | -35.2529 | -31.7752
21.6762 | -35.2534 | -31.7816
21.6823 | -35.2539 | -31.7883
21.6884 | -35.2543 | -31.7946
21.6945 | -35.2547 | -31.8011
21.7006 | -35.2552 | -31.8075
21.7068 | -35.2556 | -31.8138
21.7129 | -35.2561 | -31.8199
21.719 | -35.2565 | -31.826
21.7251 | -35.257 | -31.8324
21.7312 | -35.2574 | -31.839
21.7373 | -35.2579 | -31.8456
21.7434 | -35.2583 | -31.8522
21.7495 | -35.2588 | -31.8586
21.7556 | -35.2593 | -31.865
21.7617 | -35.2597 | -31.8717
21.7678 | -35.2601 | -31.8781
21.7739 | -35.2605 | -31.8842
21.78 | -35.261 | -31.8909
21.7861 | -35.2614 | -31.8973
21.7922 | -35.2619 | -31.9036
21.7983 | -35.2623 | -31.9097
21.8044 | -35.2627 | -31.9164
21.8105 | -35.2631 | -31.9226
21.8166 | -35.2636 | -31.929
21.8227 | -35.2641 | -31.9354
21.8288 | -35.2645 | -31.9418
21.8349 | -35.265 | -31.9484
21.841 | -35.2654 | -31.9545
21.8471 | -35.2659 | -31.9606
21.8533 | -35.2664 | -31.9668
21.8594 | -35.2669 | -31.9733
21.8655 | -35.2673 | -31.9799
21.8716 | -35.2677 | -31.9862
21.8777 | -35.2681 | -31.9926
21.8838 | -35.2686 | -31.9987
21.8899 | -35.269 | -32.0051
21.896 | -35.2694 | -32.0115
21.9021 | -35.2699 | -32.0179
21.9082 | -35.2703 | -32.0245
21.9143 | -35.2708 | -32.0306
21.9204 | -35.2712 | -32.0372
21.9265 | -35.2717 | -32.0437
21.9326 | -35.2721 | -32.05
21.9387 | -35.2726 | -32.0565
21.9448 | -35.2731 | -32.0629
21.9509 | -35.2736 | -32.0695
21.957 | -35.2741 | -32.0759
21.9631 | -35.2745 | -32.0822
21.9692 | -35.2749 | -32.0885
21.9753 | -35.2754 | -32.095
21.9814 | -35.2759 | -32.1012
21.9875 | -35.2764 | -32.1071
21.9937 | -35.2768 | -32.1135
21.9998 | -35.2772 | -32.1199
22.0059 | -35.2777 | -32.1265
22.012 | -35.2781 | -32.1328
22.0181 | -35.2786 | -32.1393
22.0242 | -35.279 | -32.1457
22.0303 | -35.2795 | -32.1522
22.0364 | -35.28 | -32.1584
22.0425 | -35.2804 | -32.1645
22.0486 | -35.2808 | -32.1709
22.0547 | -35.2813 | -32.1771
22.0608 | -35.2818 | -32.1833
22.0669 | -35.2823 | -32.1895
22.073 | -35.2827 | -32.1957
22.0791 | -35.2832 | -32.2022
22.0852 | -35.2836 | -32.2084
22.0913 | -35.284 | -32.2146
22.0974 | -35.2844 | -32.2211
22.1035 | -35.2849 | -32.2274
22.1096 | -35.2853 | -32.2339
22.1157 | -35.2857 | -32.2403
22.1218 | -35.2862 | -32.2465
22.1279 | -35.2867 | -32.2526
22.134 | -35.2872 | -32.2589
22.1402 | -35.2876 | -32.2648
22.1463 | -35.2881 | -32.271
22.1524 | -35.2885 | -32.2775
22.1585 | -35.2889 | -32.2841
22.1646 | -35.2893 | -32.29
22.1707 | -35.2898 | -32.2963
22.1768 | -35.2903 | -32.3027
22.1829 | -35.2907 | -32.309
22.189 | -35.2911 | -32.3156
22.1951 | -35.2916 | -32.3218
22.2012 | -35.2921 | -32.3279
22.2073 | -35.2926 | -32.3337
22.2134 | -35.293 | -32.3399
22.2195 | -35.2934 | -32.346
22.2256 | -35.2938 | -32.3522
22.2317 | -35.2942 | -32.3583
22.2378 | -35.2947 | -32.3648
22.2439 | -35.2951 | -32.371
22.25 | -35.2955 | -32.3773
22.2561 | -35.296 | -32.3836
22.2622 | -35.2964 | -32.3898
22.2683 | -35.2969 | -32.3961
22.2744 | -35.2973 | -32.4022
22.2806 | -35.2978 | -32.4086
22.2867 | -35.2983 | -32.4146
22.2928 | -35.2988 | -32.4209
22.2989 | -35.2993 | -32.4272
22.305 | -35.2997 | -32.4331
22.3111 | -35.3001 | -32.4394
22.3172 | -35.3006 | -32.4461
22.3233 | -35.301 | -32.4525
22.3294 | -35.3015 | -32.4592
22.3355 | -35.302 | -32.4655
22.3416 | -35.3024 | -32.4721
22.3477 | -35.3029 | -32.4783
22.3538 | -35.3033 | -32.4845
22.3599 | -35.3037 | -32.4909
22.366 | -35.3041 | -32.4973
22.3721 | -35.3046 | -32.5032
22.3782 | -35.305 | -32.5095
22.3843 | -35.3055 | -32.5158
22.3904 | -35.3059 | -32.5218
22.3965 | -35.3063 | -32.5281
22.4026 | -35.3068 | -32.5346
22.4087 | -35.3073 | -32.5407
22.4148 | -35.3077 | -32.5472
22.4209 | -35.3082 | -32.5537
22.4271 | -35.3087 | -32.5595
22.4332 | -35.3091 | -32.5658
22.4393 | -35.3096 | -32.5721
22.4454 | -35.31 | -32.5781
22.4515 | -35.3104 | -32.5843
22.4576 | -35.3109 | -32.5904
22.4637 | -35.3113 | -32.5965
22.4698 | -35.3118 | -32.6029
22.4759 | -35.3122 | -32.6089
22.482 | -35.3127 | -32.6153
22.4881 | -35.3131 | -32.6213
22.4942 | -35.3136 | -32.6279
22.5003 | -35.314 | -32.6342
22.5064 | -35.3145 | -32.6405
22.5125 | -35.315 | -32.6463
22.5186 | -35.3154 | -32.6525
22.5247 | -35.3159 | -32.6588
22.5308 | -35.3163 | -32.6651
22.5369 | -35.3168 | -32.6714
22.543 | -35.3172 | -32.6777
22.5491 | -35.3177 | -32.6839
22.5552 | -35.3181 | -32.69
22.5613 | -35.3187 | -32.6962
22.5675 | -35.3191 | -32.7022
22.5736 | -35.3195 | -32.7082
22.5797 | -35.32 | -32.7146
22.5858 | -35.3205 | -32.7202
22.5919 | -35.321 | -32.7264
22.598 | -35.3215 | -32.7329
22.6041 | -35.3219 | -32.739
22.6102 | -35.3223 | -32.7454
22.6163 | -35.3228 | -32.751
22.6224 | -35.3233 | -32.7572
22.6285 | -35.3238 | -32.7634
22.6346 | -35.3243 | -32.7695
22.6407 | -35.3247 | -32.7757
22.6468 | -35.3251 | -32.7818
22.6529 | -35.3257 | -32.7877
22.659 | -35.3261 | -32.7937
22.6651 | -35.3266 | -32.7999
22.6712 | -35.327 | -32.8063
22.6773 | -35.3274 | -32.8124
22.6834 | -35.3279 | -32.8187
22.6895 | -35.3284 | -32.8248
22.6956 | -35.3289 | -32.8308
22.7017 | -35.3293 | -32.8371
22.7079 | -35.3298 | -32.8433
22.714 | -35.3302 | -32.8494
22.7201 | -35.3306 | -32.8558
22.7262 | -35.3311 | -32.8619
22.7323 | -35.3316 | -32.8679
22.7384 | -35.332 | -32.8741
22.7445 | -35.3325 | -32.8804
22.7506 | -35.3329 | -32.8867
22.7567 | -35.3334 | -32.893
22.7628 | -35.3338 | -32.8991
22.7689 | -35.3342 | -32.9053
22.775 | -35.3347 | -32.9117
22.7811 | -35.3351 | -32.9183
22.7872 | -35.3356 | -32.9243
22.7933 | -35.3361 | -32.9303
22.7994 | -35.3365 | -32.9367
22.8055 | -35.337 | -32.9428
22.8116 | -35.3374 | -32.949
22.8177 | -35.3378 | -32.9553
22.8238 | -35.3383 | -32.9613
22.8299 | -35.3388 | -32.9673
22.836 | -35.3393 | -32.9731
22.8421 | -35.3397 | -32.9795
22.8482 | -35.3402 | -32.9855
22.8544 | -35.3407 | -32.9918
22.8605 | -35.3411 | -32.9979
22.8666 | -35.3416 | -33.0039
22.8727 | -35.3421 | -33.0097
22.8788 | -35.3425 | -33.0158
22.8849 | -35.3429 | -33.0218
22.891 | -35.3434 | -33.0279
22.8971 | -35.3438 | -33.0342
22.9032 | -35.3442 | -33.0399
22.9093 | -35.3447 | -33.0463
22.9154 | -35.3452 | -33.0524
22.9215 | -35.3456 | -33.0585
22.9276 | -35.3461 | -33.0646
22.9337 | -35.3465 | -33.0705
22.9398 | -35.3469 | -33.0767
22.9459 | -35.3474 | -33.0831
22.952 | -35.3479 | -33.0892
22.9581 | -35.3483 | -33.0952
22.9642 | -35.3488 | -33.1015
22.9703 | -35.3492 | -33.107
22.9764 | -35.3497 | -33.113
22.9825 | -35.3501 | -33.1192
22.9886 | -35.3506 | -33.1252
22.9948 | -35.351 | -33.1313
23.0009 | -35.3515 | -33.1375
23.007 | -35.352 | -33.1435
23.0131 | -35.3524 | -33.1496
23.0192 | -35.3528 | -33.1558
23.0253 | -35.3532 | -33.1619
23.0314 | -35.3537 | -33.168
23.0375 | -35.3541 | -33.1744
23.0436 | -35.3546 | -33.1801
23.0497 | -35.3551 | -33.1862
23.0558 | -35.3555 | -33.1922
23.0619 | -35.356 | -33.1981
23.068 | -35.3565 | -33.2044
23.0741 | -35.3569 | -33.2098
23.0802 | -35.3573 | -33.2155
23.0863 | -35.3578 | -33.2212
23.0924 | -35.3582 | -33.227
23.0985 | -35.3587 | -33.233
23.1046 | -35.3592 | -33.2387
23.1107 | -35.3596 | -33.2445
23.1168 | -35.3601 | -33.2506
23.1229 | -35.3605 | -33.2564
23.129 | -35.361 | -33.262
23.1351 | -35.3614 | -33.2677
23.1413 | -35.3619 | -33.2738
23.1474 | -35.3623 | -33.2795
23.1535 | -35.3628 | -33.2852
23.1596 | -35.3633 | -33.2909
23.1657 | -35.3638 | -33.2968
23.1718 | -35.3642 | -33.3026
23.1779 | -35.3646 | -33.309
23.184 | -35.3651 | -33.3144
23.1901 | -35.3656 | -33.3203
23.1962 | -35.366 | -33.3262
23.2023 | -35.3665 | -33.3322
23.2084 | -35.3669 | -33.3381
23.2145 | -35.3674 | -33.3443
23.2206 | -35.3679 | -33.3501
23.2267 | -35.3683 | -33.3563
23.2328 | -35.3687 | -33.3622
23.2389 | -35.3692 | -33.3682
23.245 | -35.3697 | -33.3744
23.2511 | -35.3701 | -33.3796
23.2572 | -35.3706 | -33.3856
23.2633 | -35.371 | -33.3916
23.2694 | -35.3714 | -33.3977
23.2755 | -35.3719 | -33.4036
23.2817 | -35.3723 | -33.4095
23.2878 | -35.3728 | -33.4152
23.2939 | -35.3732 | -33.4212
23.3 | -35.3737 | -33.4274
23.3061 | -35.3741 | -33.4334
23.3122 | -35.3745 | -33.4397
23.3183 | -35.3751 | -33.4455
23.3244 | -35.3756 | -33.4509
23.3305 | -35.3761 | -33.4568
23.3366 | -35.3765 | -33.4628
23.3427 | -35.3769 | -33.4685
23.3488 | -35.3773 | -33.4746
23.3549 | -35.3778 | -33.4808
23.361 | -35.3783 | -33.4861
23.3671 | -35.3787 | -33.4923
23.3732 | -35.3791 | -33.4983
23.3793 | -35.3796 | -33.5043
23.3854 | -35.3801 | -33.5105
23.3915 | -35.3805 | -33.5162
23.3976 | -35.381 | -33.5225
23.4037 | -35.3815 | -33.528
23.4098 | -35.3819 | -33.5339
23.4159 | -35.3824 | -33.5399
23.422 | -35.3829 | -33.5458
23.4282 | -35.3833 | -33.5513
23.4343 | -35.3838 | -33.5573
23.4404 | -35.3842 | -33.5632
23.4465 | -35.3847 | -33.5697
23.4526 | -35.3851 | -33.5757
23.4587 | -35.3856 | -33.5819
23.4648 | -35.3862 | -33.5873
23.4709 | -35.3867 | -33.5927
23.477 | -35.3871 | -33.5985
23.4831 | -35.3876 | -33.6045
23.4892 | -35.388 | -33.61
23.4953 | -35.3885 | -33.6154
23.5014 | -35.3889 | -33.6211
23.5075 | -35.3894 | -33.6275
23.5136 | -35.3899 | -33.6333
23.5197 | -35.3903 | -33.6387
23.5258 | -35.3908 | -33.6445
23.5319 | -35.3913 | -33.6499
23.538 | -35.3917 | -33.6561
23.5441 | -35.3922 | -33.6621
23.5502 | -35.3927 | -33.6681
23.5563 | -35.3931 | -33.6741
23.5624 | -35.3936 | -33.6802
23.5686 | -35.394 | -33.6851
23.5747 | -35.3945 | -33.6907
23.5808 | -35.395 | -33.6964
23.5869 | -35.3954 | -33.7022
23.593 | -35.3959 | -33.7079
23.5991 | -35.3963 | -33.7137
23.6052 | -35.3967 | -33.7198
23.6113 | -35.3972 | -33.7257
23.6174 | -35.3976 | -33.7318
23.6235 | -35.3981 | -33.7377
23.6296 | -35.3985 | -33.7436
23.6357 | -35.3989 | -33.7493
23.6418 | -35.3994 | -33.7549
23.6479 | -35.4 | -33.7608
23.654 | -35.4005 | -33.7666
23.6601 | -35.4009 | -33.7723
23.6662 | -35.4013 | -33.7779
23.6723 | -35.4018 | -33.7835
23.6784 | -35.4023 | -33.7894
23.6845 | -35.4027 | -33.795
23.6906 | -35.4032 | -33.801
23.6967 | -35.4037 | -33.8065
23.7028 | -35.4042 | -33.8126
23.7089 | -35.4046 | -33.8179
23.7151 | -35.4051 | -33.8239
23.7212 | -35.4055 | -33.83
23.7273 | -35.4059 | -33.8361
23.7334 | -35.4064 | -33.842
23.7395 | -35.4068 | -33.8477
23.7456 | -35.4072 | -33.8533
23.7517 | -35.4077 | -33.8589
23.7578 | -35.4081 | -33.8645
23.7639 | -35.4086 | -33.87
23.77 | -35.4091 | -33.8759
23.7761 | -35.4096 | -33.8819
23.7822 | -35.41 | -33.8874
23.7883 | -35.4105 | -33.8932
23.7944 | -35.411 | -33.8988
23.8005 | -35.4115 | -33.9046
23.8066 | -35.412 | -33.9105
23.8127 | -35.4124 | -33.9163
23.8188 | -35.4129 | -33.9222
23.8249 | -35.4133 | -33.928
23.831 | -35.4137 | -33.934
23.8371 | -35.4142 | -33.9396
23.8432 | -35.4146 | -33.9456
23.8493 | -35.4151 | -33.9512
23.8555 | -35.4155 | -33.9568
23.8616 | -35.416 | -33.9623
23.8677 | -35.4164 | -33.9681
23.8738 | -35.417 | -33.9735
23.8799 | -35.4174 | -33.9791
23.886 | -35.4179 | -33.9843
23.8921 | -35.4183 | -33.9901
23.8982 | -35.4188 | -33.9961
23.9043 | -35.4193 | -34.0018
23.9104 | -35.4197 | -34.0077
23.9165 | -35.4201 | -34.0132
23.9226 | -35.4206 | -34.0185
23.9287 | -35.4211 | -34.0242
23.9348 | -35.4215 | -34.0294
23.9409 | -35.4219 | -34.0353
23.947 | -35.4223 | -34.0413
23.9531 | -35.4227 | -34.0467
23.9592 | -35.4232 | -34.0525
23.9653 | -35.4236 | -34.0584
23.9714 | -35.4241 | -34.064
23.9775 | -35.4245 | -34.0697
23.9836 | -35.4249 | -34.0754
23.9897 | -35.4255 | -34.0808
23.9958 | -35.4259 | -34.0866
24.002 | -35.4263 | -34.0922
24.0081 | -35.4267 | -34.0979
24.0142 | -35.4272 | -34.1039
24.0203 | -35.4276 | -34.1094
24.0264 | -35.4281 | -34.1151
24.0325 | -35.4286 | -34.1205
24.0386 | -35.429 | -34.1264
24.0447 | -35.4295 | -34.132
24.0508 | -35.4299 | -34.1377
24.0569 | -35.4304 | -34.1433
24.063 | -35.4309 | -34.1491
24.0691 | -35.4313 | -34.1551
24.0752 | -35.4318 | -34.1606
24.0813 | -35.4323 | -34.1661
24.0874 | -35.4328 | -34.1716
24.0935 | -35.4332 | -34.1773
24.0996 | -35.4337 | -34.1826
24.1057 | -35.4341 | -34.1884
24.1118 | -35.4345 | -34.1941
24.1179 | -35.435 | -34.2001
24.124 | -35.4354 | -34.2057
24.1301 | -35.4359 | -34.2108
24.1362 | -35.4363 | -34.2163
24.1424 | -35.4369 | -34.2224
24.1485 | -35.4373 | -34.2282
24.1546 | -35.4378 | -34.2333
24.1607 | -35.4382 | -34.2389
24.1668 | -35.4387 | -34.2442
24.1729 | -35.4391 | -34.2497
24.179 | -35.4395 | -34.2557
24.1851 | -35.44 | -34.2615
24.1912 | -35.4404 | -34.267
24.1973 | -35.4408 | -34.273
24.2034 | -35.4413 | -34.2782
24.2095 | -35.4418 | -34.2843
24.2156 | -35.4422 | -34.2899
24.2217 | -35.4426 | -34.2956
24.2278 | -35.4431 | -34.3012
24.2339 | -35.4435 | -34.3068
24.24 | -35.4439 | -34.3119
24.2461 | -35.4443 | -34.3176
24.2522 | -35.4448 | -34.3231
24.2583 | -35.4453 | -34.3283
24.2644 | -35.4458 | -34.3343
24.2705 | -35.4462 | -34.3398
24.2766 | -35.4466 | -34.3451
24.2827 | -35.4471 | -34.3506
24.2889 | -35.4476 | -34.3564
24.295 | -35.448 | -34.362
24.3011 | -35.4484 | -34.3676
24.3072 | -35.4489 | -34.373
24.3133 | -35.4493 | -34.3782
24.3194 | -35.4498 | -34.3837
24.3255 | -35.4502 | -34.3892
24.3316 | -35.4506 | -34.3946
24.3377 | -35.451 | -34.3999
24.3438 | -35.4514 | -34.405
24.3499 | -35.4519 | -34.4102
24.356 | -35.4523 | -34.416
24.3621 | -35.4528 | -34.4215
24.3682 | -35.4532 | -34.4274
24.3743 | -35.4536 | -34.433
24.3804 | -35.4539 | -34.4381
24.3865 | -35.4544 | -34.4438
24.3926 | -35.4547 | -34.4487
24.3987 | -35.4552 | -34.4544
24.4048 | -35.4557 | -34.4602
24.4109 | -35.4561 | -34.4653
24.417 | -35.4566 | -34.4709
24.4231 | -35.457 | -34.4763
24.4293 | -35.4575 | -34.4817
24.4354 | -35.458 | -34.4871
24.4415 | -35.4584 | -34.4926
24.4476 | -35.4588 | -34.4981
24.4537 | -35.4593 | -34.5037
24.4598 | -35.4597 | -34.5091
24.4659 | -35.4601 | -34.5147
24.472 | -35.4606 | -34.5205
24.4781 | -35.461 | -34.5258
24.4842 | -35.4614 | -34.531
24.4903 | -35.4618 | -34.5368
24.4964 | -35.4622 | -34.5423
24.5025 | -35.4627 | -34.5482
24.5086 | -35.4631 | -34.554
24.5147 | -35.4636 | -34.5595
24.5208 | -35.4641 | -34.5645
24.5269 | -35.4645 | -34.5699
24.533 | -35.4649 | -34.5757
24.5391 | -35.4653 | -34.5813
24.5452 | -35.4658 | -34.5867
24.5513 | -35.4661 | -34.5918
24.5574 | -35.4666 | -34.5972
24.5635 | -35.467 | -34.6027
24.5696 | -35.4674 | -34.6084
24.5758 | -35.4679 | -34.614
24.5819 | -35.4683 | -34.6196
24.588 | -35.4688 | -34.6256
24.5941 | -35.4693 | -34.6315
24.6002 | -35.4697 | -34.6374
24.6063 | -35.4701 | -34.6432
24.6124 | -35.4706 | -34.6489
24.6185 | -35.471 | -34.6547
24.6246 | -35.4714 | -34.6597
24.6307 | -35.4718 | -34.6651
24.6368 | -35.4723 | -34.6706
24.6429 | -35.4727 | -34.676
24.649 | -35.4732 | -34.682
24.6551 | -35.4736 | -34.6874
24.6612 | -35.474 | -34.6922
24.6673 | -35.4745 | -34.6978
24.6734 | -35.475 | -34.7033
24.6795 | -35.4754 | -34.7089
24.6856 | -35.4758 | -34.7144
24.6917 | -35.4762 | -34.72
24.6978 | -35.4766 | -34.7252
24.7039 | -35.477 | -34.7308
24.71 | -35.4774 | -34.7362
24.7162 | -35.4778 | -34.7417
24.7223 | -35.4783 | -34.7472
24.7284 | -35.4788 | -34.7522
24.7345 | -35.4792 | -34.7576
24.7406 | -35.4796 | -34.7631
24.7467 | -35.48 | -34.7688
24.7528 | -35.4805 | -34.7745
24.7589 | -35.4809 | -34.7801
24.765 | -35.4814 | -34.7854
24.7711 | -35.4818 | -34.791
24.7772 | -35.4823 | -34.7967
24.7833 | -35.4827 | -34.8023
24.7894 | -35.4832 | -34.8076
24.7955 | -35.4836 | -34.8134
24.8016 | -35.4841 | -34.8185
24.8077 | -35.4845 | -34.8242
24.8138 | -35.4849 | -34.8297
24.8199 | -35.4854 | -34.8352
24.826 | -35.4858 | -34.8404
24.8321 | -35.4862 | -34.8461
24.8382 | -35.4866 | -34.8512
24.8443 | -35.487 | -34.8567
24.8504 | -35.4874 | -34.862
24.8565 | -35.4878 | -34.8671
24.8627 | -35.4883 | -34.8721
24.8688 | -35.4886 | -34.8781
24.8749 | -35.489 | -34.8836
24.881 | -35.4894 | -34.8886
24.8871 | -35.4898 | -34.8942
24.8932 | -35.4903 | -34.8998
24.8993 | -35.4907 | -34.905
24.9054 | -35.4911 | -34.9106
24.9115 | -35.4916 | -34.9162
24.9176 | -35.492 | -34.9216
24.9237 | -35.4925 | -34.927
24.9298 | -35.4929 | -34.9324
24.9359 | -35.4933 | -34.9379
24.942 | -35.4936 | -34.9431
24.9481 | -35.4941 | -34.9482
24.9542 | -35.4945 | -34.9538
24.9603 | -35.4949 | -34.9594
24.9664 | -35.4954 | -34.9649
24.9725 | -35.4958 | -34.9706
24.9786 | -35.4962 | -34.9761
24.9847 | -35.4967 | -34.9815
24.9908 | -35.4971 | -34.9871
24.9969 | -35.4975 | -34.9925
25.0031 | -35.498 | -34.9978
25.0092 | -35.4984 | -35.003
25.0153 | -35.4988 | -35.0083
25.0214 | -35.4993 | -35.014
25.0275 | -35.4997 | -35.0193
25.0336 | -35.5001 | -35.0249
25.0397 | -35.5005 | -35.0304
25.0458 | -35.5009 | -35.0355
25.0519 | -35.5014 | -35.0408
25.058 | -35.5018 | -35.0462
25.0641 | -35.5022 | -35.0513
25.0702 | -35.5026 | -35.0565
25.0763 | -35.503 | -35.0621
25.0824 | -35.5034 | -35.0669
25.0885 | -35.5039 | -35.0723
25.0946 | -35.5043 | -35.0775
25.1007 | -35.5046 | -35.0829
25.1068 | -35.5051 | -35.0887
25.1129 | -35.5055 | -35.094
25.119 | -35.5059 | -35.0994
25.1251 | -35.5063 | -35.1047
25.1312 | -35.5068 | -35.11
25.1373 | -35.5072 | -35.1156
25.1435 | -35.5076 | -35.1209
25.1496 | -35.5081 | -35.1263
25.1557 | -35.5086 | -35.1319
25.1618 | -35.509 | -35.1375
25.1679 | -35.5094 | -35.1435
25.174 | -35.5098 | -35.1485
25.1801 | -35.5102 | -35.1541
25.1862 | -35.5106 | -35.1589
25.1923 | -35.5111 | -35.1645
25.1984 | -35.5115 | -35.17
25.2045 | -35.5119 | -35.1755
25.2106 | -35.5123 | -35.1811
25.2167 | -35.5127 | -35.1863
25.2228 | -35.513 | -35.1914
25.2289 | -35.5134 | -35.1969
25.235 | -35.5139 | -35.2021
25.2411 | -35.5143 | -35.2076
25.2472 | -35.5147 | -35.2129
25.2533 | -35.515 | -35.2185
25.2594 | -35.5154 | -35.2235
25.2655 | -35.5158 | -35.2286
25.2716 | -35.5163 | -35.2338
25.2777 | -35.5167 | -35.2392
25.2838 | -35.5171 | -35.2449
25.29 | -35.5175 | -35.2503
25.2961 | -35.5179 | -35.2554
25.3022 | -35.5183 | -35.261
25.3083 | -35.5187 | -35.2663
25.3144 | -35.5191 | -35.2717
25.3205 | -35.5196 | -35.277
25.3266 | -35.5199 | -35.2821
25.3327 | -35.5204 | -35.2876
25.3388 | -35.5208 | -35.293
25.3449 | -35.5212 | -35.2981
25.351 | -35.5217 | -35.3033
25.3571 | -35.5221 | -35.3087
25.3632 | -35.5224 | -35.3138
25.3693 | -35.5229 | -35.3194
25.3754 | -35.5233 | -35.3248
25.3815 | -35.5237 | -35.3301
25.3876 | -35.5241 | -35.3354
25.3937 | -35.5245 | -35.3406
25.3998 | -35.5249 | -35.346
25.4059 | -35.5253 | -35.3513
25.412 | -35.5258 | -35.3567
25.4181 | -35.5262 | -35.3618
25.4242 | -35.5267 | -35.3671
25.4304 | -35.5271 | -35.3725
25.4365 | -35.5275 | -35.3776
25.4426 | -35.5279 | -35.383
25.4487 | -35.5283 | -35.3884
25.4548 | -35.5287 | -35.3938
25.4609 | -35.5291 | -35.3994
25.467 | -35.5295 | -35.4042
25.4731 | -35.5299 | -35.4093
25.4792 | -35.5303 | -35.4149
25.4853 | -35.5307 | -35.4203
25.4914 | -35.5312 | -35.4258
25.4975 | -35.5316 | -35.4313
25.5036 | -35.532 | -35.4362
25.5097 | -35.5324 | -35.4416
25.5158 | -35.5328 | -35.4466
25.5219 | -35.5332 | -35.4518
25.528 | -35.5336 | -35.457
25.5341 | -35.5341 | -35.4626
25.5402 | -35.5345 | -35.468
25.5463 | -35.5348 | -35.4735
25.5524 | -35.5352 | -35.4786
25.5585 | -35.5357 | -35.4838
25.5646 | -35.536 | -35.4889
25.5707 | -35.5364 | -35.4941
25.5769 | -35.5368 | -35.4994
25.583 | -35.5371 | -35.5043
25.5891 | -35.5375 | -35.51
25.5952 | -35.538 | -35.5151
25.6013 | -35.5384 | -35.5203
25.6074 | -35.5388 | -35.5254
25.6135 | -35.5392 | -35.5307
25.6196 | -35.5396 | -35.5361
25.6257 | -35.54 | -35.5417
25.6318 | -35.5403 | -35.5465
25.6379 | -35.5407 | -35.5518
25.644 | -35.5411 | -35.5567
25.6501 | -35.5415 | -35.5618
25.6562 | -35.5419 | -35.5671
25.6623 | -35.5423 | -35.5723
25.6684 | -35.5428 | -35.5772
25.6745 | -35.5432 | -35.5826
25.6806 | -35.5436 | -35.5877
25.6867 | -35.5439 | -35.5928
25.6928 | -35.5443 | -35.5982
25.6989 | -35.5447 | -35.6038
25.705 | -35.5451 | -35.609
25.7111 | -35.5455 | -35.6146
25.7173 | -35.5458 | -35.6195
25.7234 | -35.5462 | -35.625
25.7295 | -35.5467 | -35.6301
25.7356 | -35.5471 | -35.6358
25.7417 | -35.5475 | -35.6411
25.7478 | -35.5479 | -35.646
25.7539 | -35.5483 | -35.6512
25.76 | -35.5487 | -35.6567
25.7661 | -35.5491 | -35.6619
25.7722 | -35.5495 | -35.6673
25.7783 | -35.5498 | -35.6725
25.7844 | -35.5502 | -35.6774
25.7905 | -35.5506 | -35.6829
25.7966 | -35.551 | -35.6881
25.8027 | -35.5514 | -35.6935
25.8088 | -35.5518 | -35.6988
25.8149 | -35.5522 | -35.7039
25.821 | -35.5525 | -35.7091
25.8271 | -35.5529 | -35.7144
25.8332 | -35.5533 | -35.7197
25.8393 | -35.5537 | -35.7252
25.8454 | -35.5541 | -35.7302
25.8515 | -35.5545 | -35.7354
25.8576 | -35.5549 | -35.7408
25.8638 | -35.5553 | -35.7463
25.8699 | -35.5557 | -35.7517
25.876 | -35.5561 | -35.7571
25.8821 | -35.5565 | -35.7623
25.8882 | -35.5569 | -35.7674
25.8943 | -35.5573 | -35.7725
25.9004 | -35.5577 | -35.778
25.9065 | -35.5581 | -35.7834
25.9126 | -35.5585 | -35.7889
25.9187 | -35.5589 | -35.7937
25.9248 | -35.5593 | -35.7987
25.9309 | -35.5597 | -35.8038
25.937 | -35.5601 | -35.8091
25.9431 | -35.5605 | -35.8145
25.9492 | -35.5608 | -35.8194
25.9553 | -35.5612 | -35.8248
25.9614 | -35.5616 | -35.83
25.9675 | -35.5619 | -35.8351
25.9736 | -35.5623 | -35.8402
25.9797 | -35.5627 | -35.8455
25.9858 | -35.5631 | -35.8507
25.9919 | -35.5635 | -35.8562
25.998 | -35.5638 | -35.8613
26.0042 | -35.5642 | -35.8664
26.0103 | -35.5646 | -35.8714
26.0164 | -35.5649 | -35.8764
26.0225 | -35.5653 | -35.8814
26.0286 | -35.5657 | -35.8869
26.0347 | -35.5661 | -35.892
26.0408 | -35.5665 | -35.8972
26.0469 | -35.5668 | -35.9021
26.053 | -35.5672 | -35.9075
26.0591 | -35.5676 | -35.9127
26.0652 | -35.5679 | -35.9182
26.0713 | -35.5683 | -35.9233
26.0774 | -35.5687 | -35.9284
26.0835 | -35.5691 | -35.9338
26.0896 | -35.5695 | -35.939
26.0957 | -35.5699 | -35.944
26.1018 | -35.5703 | -35.9488
26.1079 | -35.5707 | -35.9541
26.114 | -35.5711 | -35.9592
26.1201 | -35.5715 | -35.9643
26.1262 | -35.5719 | -35.9694
26.1323 | -35.5722 | -35.9747
26.1384 | -35.5726 | -35.9798
26.1445 | -35.573 | -35.9852
26.1507 | -35.5733 | -35.9906
26.1568 | -35.5737 | -35.9956
26.1629 | -35.5741 | -36.0007
26.169 | -35.5745 | -36.006
26.1751 | -35.5749 | -36.011
26.1812 | -35.5752 | -36.0166
26.1873 | -35.5756 | -36.0223
26.1934 | -35.576 | -36.0275
26.1995 | -35.5764 | -36.0331
26.2056 | -35.5767 | -36.0381
26.2117 | -35.5771 | -36.0434
26.2178 | -35.5774 | -36.0486
26.2239 | -35.5778 | -36.054
26.23 | -35.5781 | -36.0591
26.2361 | -35.5785 | -36.0644
26.2422 | -35.5789 | -36.0694
26.2483 | -35.5793 | -36.0747
26.2544 | -35.5796 | -36.0802
26.2605 | -35.58 | -36.0857
26.2666 | -35.5804 | -36.091
26.2727 | -35.5807 | -36.0963
26.2788 | -35.5811 | -36.101
26.2849 | -35.5814 | -36.106
26.2911 | -35.5818 | -36.1111
26.2972 | -35.5822 | -36.1165
26.3033 | -35.5826 | -36.1217
26.3094 | -35.5829 | -36.1265
26.3155 | -35.5833 | -36.1312
26.3216 | -35.5836 | -36.1365
26.3277 | -35.5839 | -36.1416
26.3338 | -35.5843 | -36.1465
26.3399 | -35.5847 | -36.1513
26.346 | -35.5849 | -36.1564
26.3521 | -35.5853 | -36.1617
26.3582 | -35.5857 | -36.1667
26.3643 | -35.5861 | -36.1718
26.3704 | -35.5865 | -36.1771
26.3765 | -35.5868 | -36.1826
26.3826 | -35.5871 | -36.1878
26.3887 | -35.5875 | -36.1933
26.3948 | -35.5878 | -36.1988
26.4009 | -35.5882 | -36.2042
26.407 | -35.5885 | -36.2097
26.4131 | -35.5889 | -36.2147
26.4192 | -35.5892 | -36.2198
26.4253 | -35.5896 | -36.2252
26.4314 | -35.5899 | -36.2305
26.4376 | -35.5903 | -36.2356
26.4437 | -35.5907 | -36.2412
26.4498 | -35.591 | -36.2463
26.4559 | -35.5914 | -36.2513
26.462 | -35.5917 | -36.2562
26.4681 | -35.592 | -36.2615
26.4742 | -35.5924 | -36.2665
26.4803 | -35.5927 | -36.2716
26.4864 | -35.5931 | -36.2764
26.4925 | -35.5934 | -36.2816
26.4986 | -35.5939 | -36.2868
26.5047 | -35.5942 | -36.292
26.5108 | -35.5946 | -36.297
26.5169 | -35.5949 | -36.3023
26.523 | -35.5953 | -36.3072
26.5291 | -35.5957 | -36.3125
26.5352 | -35.5961 | -36.3176
26.5413 | -35.5965 | -36.3228
26.5474 | -35.5968 | -36.3281
26.5535 | -35.5972 | -36.3333
26.5596 | -35.5975 | -36.3383
26.5657 | -35.5979 | -36.3437
26.5718 | -35.5982 | -36.3485
26.578 | -35.5985 | -36.3538
26.5841 | -35.5989 | -36.3589
26.5902 | -35.5993 | -36.3635
26.5963 | -35.5996 | -36.3683
26.6024 | -35.6 | -36.3733
26.6085 | -35.6003 | -36.3789
26.6146 | -35.6007 | -36.384
26.6207 | -35.601 | -36.3891
26.6268 | -35.6013 | -36.3944
26.6329 | -35.6016 | -36.3992
26.639 | -35.602 | -36.4042
26.6451 | -35.6023 | -36.4092
26.6512 | -35.6027 | -36.4149
26.6573 | -35.6031 | -36.4199
26.6634 | -35.6034 | -36.4248
26.6695 | -35.6038 | -36.43
26.6756 | -35.6041 | -36.435
26.6817 | -35.6044 | -36.4399
26.6878 | -35.6047 | -36.4452
26.6939 | -35.6051 | -36.4503
26.7 | -35.6054 | -36.455
26.7061 | -35.6057 | -36.4601
26.7122 | -35.6061 | -36.4648
26.7183 | -35.6065 | -36.4698
26.7245 | -35.6068 | -36.4751
26.7306 | -35.6071 | -36.4807
26.7367 | -35.6074 | -36.4859
26.7428 | -35.6078 | -36.4913
26.7489 | -35.6082 | -36.4969
26.755 | -35.6085 | -36.5018
26.7611 | -35.6089 | -36.5073
26.7672 | -35.6092 | -36.5124
26.7733 | -35.6096 | -36.5176
26.7794 | -35.6099 | -36.5227
26.7855 | -35.6102 | -36.528
26.7916 | -35.6106 | -36.5333
26.7977 | -35.6109 | -36.5386
26.8038 | -35.6112 | -36.5437
26.8099 | -35.6115 | -36.5492
26.816 | -35.6119 | -36.5544
26.8221 | -35.6122 | -36.5595
26.8282 | -35.6125 | -36.5646
26.8343 | -35.6128 | -36.57
26.8404 | -35.6131 | -36.5748
26.8465 | -35.6135 | -36.5799
26.8526 | -35.6138 | -36.5851
26.8587 | -35.6141 | -36.5899
26.8649 | -35.6145 | -36.5951
26.871 | -35.6148 | -36.6002
26.8771 | -35.6151 | -36.6056
26.8832 | -35.6154 | -36.6109
26.8893 | -35.6157 | -36.6163
26.8954 | -35.616 | -36.6213
26.9015 | -35.6164 | -36.6266
26.9076 | -35.6167 | -36.632
26.9137 | -35.617 | -36.6375
26.9198 | -35.6173 | -36.6428
26.9259 | -35.6176 | -36.6481
26.932 | -35.618 | -36.6531
26.9381 | -35.6184 | -36.6584
26.9442 | -35.6187 | -36.6633
26.9503 | -35.619 | -36.6681
26.9564 | -35.6193 | -36.673
26.9625 | -35.6196 | -36.6782
26.9686 | -35.62 | -36.6833
26.9747 | -35.6203 | -36.688
26.9808 | -35.6206 | -36.6929
26.9869 | -35.6209 | -36.6984
26.993 | -35.6212 | -36.7038
26.9991 | -35.6216 | -36.7088
27.0052 | -35.6219 | -36.714
27.0114 | -35.6222 | -36.7193
27.0175 | -35.6224 | -36.7241
27.0236 | -35.6228 | -36.7293
27.0297 | -35.6231 | -36.7343
27.0358 | -35.6234 | -36.7397
27.0419 | -35.6237 | -36.7453
27.048 | -35.6241 | -36.7507
27.0541 | -35.6244 | -36.7557
27.0602 | -35.6247 | -36.7612
27.0663 | -35.6251 | -36.7661
27.0724 | -35.6255 | -36.7715
27.0785 | -35.6258 | -36.7768
27.0846 | -35.6261 | -36.7816
27.0907 | -35.6264 | -36.7869
27.0968 | -35.6267 | -36.7919
27.1029 | -35.6271 | -36.7973
27.109 | -35.6274 | -36.8023
27.1151 | -35.6277 | -36.8075
27.1212 | -35.6281 | -36.8126
27.1273 | -35.6284 | -36.8178
27.1334 | -35.6287 | -36.8232
27.1395 | -35.6291 | -36.8283
27.1456 | -35.6294 | -36.8333
27.1518 | -35.6297 | -36.8385
27.1579 | -35.63 | -36.8433
27.164 | -35.6303 | -36.8488
27.1701 | -35.6307 | -36.8543
27.1762 | -35.6311 | -36.8597
27.1823 | -35.6313 | -36.8648
27.1884 | -35.6316 | -36.8699
27.1945 | -35.6318 | -36.8749
27.2006 | -35.6321 | -36.88
27.2067 | -35.6325 | -36.8849
27.2128 | -35.6328 | -36.8904
27.2189 | -35.6331 | -36.895
27.225 | -35.6334 | -36.9004
27.2311 | -35.6336 | -36.9057
27.2372 | -35.6339 | -36.9111
27.2433 | -35.6343 | -36.916
27.2494 | -35.6346 | -36.921
27.2555 | -35.635 | -36.926
27.2616 | -35.6353 | -36.9307
27.2677 | -35.6356 | -36.936
27.2738 | -35.6359 | -36.9408
27.2799 | -35.6362 | -36.9457
27.286 | -35.6366 | -36.9506
27.2921 | -35.6369 | -36.956
27.2983 | -35.6372 | -36.9613
27.3044 | -35.6375 | -36.9668
27.3105 | -35.6378 | -36.972
27.3166 | -35.6381 | -36.9778
27.3227 | -35.6384 | -36.983
27.3288 | -35.6388 | -36.9883
27.3349 | -35.639 | -36.9936
27.341 | -35.6393 | -36.9986
27.3471 | -35.6396 | -37.0042
27.3532 | -35.64 | -37.0091
27.3593 | -35.6403 | -37.0148
27.3654 | -35.6406 | -37.02
27.3715 | -35.6408 | -37.0249
27.3776 | -35.6411 | -37.0302
27.3837 | -35.6414 | -37.0358
27.3898 | -35.6417 | -37.0414
27.3959 | -35.6419 | -37.0463
27.402 | -35.6421 | -37.0515
27.4081 | -35.6424 | -37.0569
27.4142 | -35.6427 | -37.0623
27.4203 | -35.643 | -37.0678
27.4264 | -35.6432 | -37.0727
27.4325 | -35.6436 | -37.0777
27.4387 | -35.6438 | -37.0833
27.4448 | -35.6442 | -37.0886
27.4509 | -35.6444 | -37.094
27.457 | -35.6448 | -37.0991
27.4631 | -35.6451 | -37.1043
27.4692 | -35.6454 | -37.1098
27.4753 | -35.6457 | -37.1142
27.4814 | -35.646 | -37.1191
27.4875 | -35.6463 | -37.1243
27.4936 | -35.6465 | -37.1294
27.4997 | -35.6468 | -37.1343
27.5058 | -35.6471 | -37.1394
27.5119 | -35.6473 | -37.1439
27.518 | -35.6476 | -37.1491
27.5241 | -35.6479 | -37.1545
27.5302 | -35.6482 | -37.1594
27.5363 | -35.6484 | -37.1644
27.5424 | -35.6487 | -37.1692
27.5485 | -35.649 | -37.1743
27.5546 | -35.6494 | -37.1793
27.5607 | -35.6496 | -37.184
27.5668 | -35.6499 | -37.1897
27.5729 | -35.6502 | -37.1951
27.5791 | -35.6504 | -37.2004
27.5852 | -35.6507 | -37.2053
27.5913 | -35.6509 | -37.2104
27.5974 | -35.6512 | -37.2152
27.6035 | -35.6514 | -37.2205
27.6096 | -35.6517 | -37.2252
27.6157 | -35.6519 | -37.2295
27.6218 | -35.6522 | -37.2349
27.6279 | -35.6524 | -37.2394
27.634 | -35.6527 | -37.2446
27.6401 | -35.6529 | -37.2495
27.6462 | -35.6531 | -37.2542
27.6523 | -35.6535 | -37.2577
27.6584 | -35.6538 | -37.2624
27.6645 | -35.6541 | -37.2678
27.6706 | -35.6544 | -37.2724
27.6767 | -35.6547 | -37.2771
27.6828 | -35.655 | -37.2824
27.6889 | -35.6553 | -37.2874
27.695 | -35.6555 | -37.2922
27.7011 | -35.6557 | -37.2971
27.7072 | -35.6561 | -37.3024
27.7133 | -35.6565 | -37.3075
27.7194 | -35.6568 | -37.3127
27.7256 | -35.6571 | -37.3178
27.7317 | -35.6574 | -37.3231
27.7378 | -35.6577 | -37.328
27.7439 | -35.658 | -37.3329
27.75 | -35.6583 | -37.3376
27.7561 | -35.6585 | -37.3435
27.7622 | -35.6587 | -37.3487
27.7683 | -35.6591 | -37.3538
27.7744 | -35.6594 | -37.3585
27.7805 | -35.6597 | -37.3631
27.7866 | -35.66 | -37.3681
27.7927 | -35.6604 | -37.373
27.7988 | -35.6607 | -37.3783
27.8049 | -35.6609 | -37.3828
27.811 | -35.6613 | -37.3884
27.8171 | -35.6616 | -37.3935
27.8232 | -35.6618 | -37.3977
27.8293 | -35.6621 | -37.4029
27.8354 | -35.6624 | -37.4079
27.8415 | -35.6627 | -37.4126
27.8476 | -35.6629 | -37.4175
27.8537 | -35.6632 | -37.4221
27.8598 | -35.6634 | -37.4273
27.866 | -35.6637 | -37.4323
27.8721 | -35.6639 | -37.4375
27.8782 | -35.6642 | -37.4426
27.8843 | -35.6645 | -37.4475
27.8904 | -35.6647 | -37.4525
27.8965 | -35.665 | -37.4575
27.9026 | -35.6653 | -37.4621
27.9087 | -35.6656 | -37.4677
27.9148 | -35.6658 | -37.4724
27.9209 | -35.6661 | -37.4771
27.927 | -35.6664 | -37.4826
27.9331 | -35.6667 | -37.4878
27.9392 | -35.6671 | -37.493
27.9453 | -35.6674 | -37.4985
27.9514 | -35.6677 | -37.503
27.9575 | -35.668 | -37.5082
27.9636 | -35.6683 | -37.5142
27.9697 | -35.6684 | -37.5194
27.9758 | -35.6687 | -37.5248
27.9819 | -35.669 | -37.5298
27.988 | -35.6693 | -37.5355
27.9941 | -35.6696 | -37.5398
28.0002 | -35.6698 | -37.5447
28.0063 | -35.67 | -37.5498
28.0125 | -35.6704 | -37.5548
28.0186 | -35.6708 | -37.5594
28.0247 | -35.671 | -37.5642
28.0308 | -35.6713 | -37.5693
28.0369 | -35.6715 | -37.5737
28.043 | -35.6717 | -37.5793
28.0491 | -35.672 | -37.5842
28.0552 | -35.6724 | -37.5896
28.0613 | -35.6726 | -37.5942
28.0674 | -35.673 | -37.5991
28.0735 | -35.6733 | -37.6037

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.3107 | -41.7142
21.4908 | -35.4958 | -34.8035
16.4948 | -34.8662 | -27.5103
12.6602 | -34.5096 | -21.2782
9.7171 | -34.2942 | -16.4135
7.4582 | -34.1453 | -12.4687
5.7244 | -34.0336 | -9.2685
4.3936 | -33.9644 | -6.6894
3.3722 | -33.9523 | -4.6571
2.5883 | -33.9551 | -3.1121
1.9866 | -33.9535 | -1.8825
1.5248 | -33.9534 | -0.89177
1.1703 | -33.9586 | -0.24699
1 | -33.9568 | 0.072277
0.89822 | -33.9571 | 0.34493
0.68941 | -33.9568 | 0.9224
0.52919 | -33.9523 | 1.3739
0.40612 | -33.9319 | 1.9776
0.31171 | -33.9205 | 2.4267
0.23928 | -33.9416 | 3.1459
0.18358 | -33.9226 | 3.5831
0.14091 | -33.9531 | 4.4641
0.10819 | -33.9526 | 5.6357
0.083017 | -33.9876 | 7.1194
0.063705 | -34.0345 | 8.9958
0.048923 | -34.1199 | 11.5015
0.037479 | -34.2377 | 14.6979
0.028801 | -34.4118 | 18.7253
0.022078 | -34.7275 | 24.0621
0.017023 | -35.2178 | 30.4467
0.013018 | -36.0356 | 38.1134
0.010014 | -37.1325 | 46.0318

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003727 | -10.1918
0.007454 | -28.2132
0.011181 | -35.9232
0.014908 | -40.8435
0.018635 | -42.2389
0.022362 | -42.94
0.026089 | -44.4041
0.029816 | -46.8638
0.033543 | -48.7653
0.03727 | -49.5914
0.040997 | -51.3006
0.044724 | -52.9564
0.048451 | -53.7239
0.052178 | -54.5497
0.055905 | -55.937
0.059632 | -57.3528
0.063359 | -57.1
0.067086 | -57.5655
0.070813 | -57.781
0.07454 | -59.9508
0.078267 | -60.8861
0.081994 | -59.9855
0.085721 | -60.0491
0.089448 | -61.6284
0.093175 | -61.7423
0.096902 | -62.2205
0.10063 | -62.7905
0.10436 | -62.2222
0.10808 | -62.5806
0.11181 | -63.3043
0.11554 | -64.1241
0.11926 | -64.3176
0.12299 | -64.9716
0.12672 | -65.5312
0.13045 | -65.3444
0.13417 | -65.7954
0.1379 | -66.4708
0.14163 | -65.7043
0.14535 | -65.8358
0.14908 | -67.0823
0.15281 | -67.285
0.15653 | -67.2208
0.16026 | -67.5481
0.16399 | -67.6466
0.16772 | -68.2621
0.17144 | -68.4294
0.17517 | -67.6384
0.1789 | -68.2536
0.18262 | -69.5609
0.18635 | -69.9733
0.19008 | -69.4608
0.1938 | -69.7669
0.19753 | -70.4892
0.20126 | -69.0564
0.20499 | -69.6281
0.20871 | -70.9831
0.21244 | -71.0996
0.21617 | -71.1954
0.21989 | -71.3815
0.22362 | -72.2576
0.22735 | -72.4528
0.23107 | -72.4068
0.2348 | -71.4077
0.23853 | -72.6165
0.24225 | -73.4191
0.24598 | -72.698
0.24971 | -73.4609
0.25344 | -73.7613
0.25716 | -73.5887
0.26089 | -72.9156
0.26462 | -74.0427
0.26834 | -75.0808
0.27207 | -74.4572
0.2758 | -74.3636
0.27953 | -74.4875
0.28325 | -74.3837
0.28698 | -74.3301
0.29071 | -74.7476
0.29443 | -75.8341
0.29816 | -75.5512
0.30189 | -75.5939
0.30561 | -76.5929
0.30934 | -76.2598
0.31307 | -75.6792
0.31679 | -75.2132
0.32052 | -75.6091
0.32425 | -77.1036
0.32798 | -77.8878
0.3317 | -77.6816
0.33543 | -77.121
0.33916 | -76.158
0.34288 | -76.7084
0.34661 | -76.5225
0.35034 | -75.6504
0.35407 | -76.5264
0.35779 | -76.4564
0.36152 | -76.7525
0.36525 | -77.0453
0.36897 | -76.6196
0.3727 | -76.9265
0.37643 | -77.1239
0.38015 | -77.3936
0.38388 | -77.1587
0.38761 | -77.7663
0.39133 | -77.7675
0.39506 | -78.246
0.39879 | -79.9601
0.40252 | -78.8735
0.40624 | -78.8705
0.40997 | -78.9721
0.4137 | -79.1269
0.41742 | -78.4787
0.42115 | -78.2398
0.42488 | -79.2661
0.42861 | -78.9349
0.43233 | -78.874
0.43606 | -79.1576
0.43979 | -78.724
0.44351 | -78.2302
0.44724 | -77.6228
0.45097 | -79.3568
0.45469 | -79.7452
0.45842 | -79.2296
0.46215 | -79.0897
0.46587 | -78.9003
0.4696 | -79.6334
0.47333 | -79.7916
0.47706 | -79.9547
0.48078 | -80.3781
0.48451 | -80.4515
0.48824 | -79.9092
0.49196 | -79.9298
0.49569 | -80.0255
0.49942 | -79.9933
0.50314 | -80.5033
0.50687 | -80.8227
0.5106 | -81.3631
0.51433 | -80.9204
0.51805 | -80.4757
0.52178 | -81.0921
0.52551 | -81.1773
0.52923 | -81.5423
0.53296 | -80.7697
0.53669 | -81.2954
0.54041 | -81.4766
0.54414 | -81.4772
0.54787 | -81.2121
0.5516 | -80.8267
0.55532 | -80.7092
0.55905 | -80.8634
0.56278 | -81.214
0.5665 | -81.8423
0.57023 | -81.9189
0.57396 | -81.8961
0.57769 | -82.1446
0.58141 | -82.1748
0.58514 | -81.8656
0.58887 | -81.801
0.59259 | -82.3618
0.59632 | -82.8525
0.60005 | -82.9303
0.60377 | -82.8998
0.6075 | -82.565
0.61123 | -82.4416
0.61496 | -82.2995
0.61868 | -82.0101
0.62241 | -82.2757
0.62614 | -82.8825
0.62986 | -82.8343
0.63359 | -82.3678
0.63732 | -82.1686
0.64104 | -82.6875
0.64477 | -83.461
0.6485 | -84.029
0.65223 | -84.0671
0.65595 | -83.5286
0.65968 | -82.8718
0.66341 | -82.8766
0.66713 | -83.0242
0.67086 | -82.8779
0.67459 | -83.0334
0.67831 | -83.7879
0.68204 | -84.4838
0.68577 | -84.4687
0.68949 | -84.2784
0.69322 | -83.9663
0.69695 | -83.8622
0.70068 | -84.4019
0.7044 | -84.6862
0.70813 | -84.2594
0.71186 | -83.9064
0.71558 | -83.8964
0.71931 | -83.6087
0.72304 | -83.0766
0.72676 | -82.9364
0.73049 | -83.4016
0.73422 | -84.1356
0.73794 | -84.8131
0.74167 | -84.8177
0.7454 | -84.5607
0.74913 | -84.6052
0.75285 | -84.7511
0.75658 | -84.7337
0.76031 | -84.5052
0.76403 | -84.3559
0.76776 | -84.8875
0.77149 | -85.548
0.77521 | -85.3195
0.77894 | -84.7477
0.78267 | -85.1236
0.7864 | -85.9179
0.79012 | -85.8366
0.79385 | -85.1802
0.79758 | -84.9734
0.8013 | -84.9506
0.80503 | -85.0439
0.80876 | -85.2796
0.81249 | -85.5577
0.81621 | -85.8527
0.81994 | -86.2809
0.82367 | -86.4203
0.82739 | -85.8419
0.83112 | -85.1684
0.83485 | -84.946
0.83857 | -85.0152
0.8423 | -85.3603
0.84603 | -85.5695
0.84976 | -85.633
0.85348 | -85.6878
0.85721 | -85.6622
0.86094 | -85.9993
0.86466 | -86.4143
0.86839 | -86.1965
0.87212 | -85.8057
0.87584 | -85.8781
0.87957 | -86.0791
0.8833 | -86.3801
0.88702 | -86.6391
0.89075 | -86.7776
0.89448 | -86.4974
0.89821 | -85.8232
0.90193 | -85.8384
0.90566 | -86.5295
0.90939 | -86.9517
0.91311 | -86.8377
0.91684 | -86.4867
0.92057 | -86.2802
0.92429 | -86.0902
0.92802 | -85.9502
0.93175 | -86.1224
0.93548 | -86.4818
0.9392 | -87.0088
0.94293 | -87.5649
0.94666 | -87.6995
0.95038 | -87.2075
0.95411 | -86.3923
0.95784 | -86.0184
0.96157 | -85.9894
0.96529 | -85.9502
0.96902 | -86.2176
0.97275 | -86.3406
0.97647 | -86.2946
0.9802 | -86.5553
0.98393 | -86.9262
0.98765 | -86.7798
0.99138 | -86.5629
0.99511 | -86.9018
0.99884 | -87.2679
1.0026 | -87.6353
1.0063 | -88.0121
1.01 | -87.7527
1.0137 | -87.3847
1.0175 | -87.1977
1.0212 | -86.945
1.0249 | -87.1069
1.0287 | -87.5751
1.0324 | -87.9475
1.0361 | -87.9458
1.0398 | -88.0328
1.0436 | -88.232
1.0473 | -88.0852
1.051 | -87.875
1.0547 | -88.2576
1.0585 | -88.9657
1.0622 | -89.1214
1.0659 | -88.8143
1.0696 | -88.0576
1.0734 | -87.7944
1.0771 | -87.7611
1.0808 | -87.927
1.0846 | -88.103
1.0883 | -88.242
1.092 | -88.4639
1.0957 | -88.721
1.0995 | -88.8359
1.1032 | -88.9457
1.1069 | -89.2139
1.1106 | -89.363
1.1144 | -89.1625
1.1181 | -88.8388
1.1218 | -88.595
1.1256 | -88.5066
1.1293 | -88.6074
1.133 | -88.7836
1.1367 | -88.9199
1.1405 | -89.058
1.1442 | -89.195
1.1479 | -89.2152
1.1516 | -89.0614
1.1554 | -88.792
1.1591 | -88.5631
1.1628 | -88.5238
1.1665 | -88.5767
1.1703 | -88.5181
1.174 | -88.491
1.1777 | -88.7128
1.1815 | -89.0876
1.1852 | -89.273
1.1889 | -89.0169
1.1926 | -88.5605
1.1964 | -88.4098
1.2001 | -88.8021
1.2038 | -89.5354
1.2075 | -89.993
1.2113 | -89.8934
1.215 | -89.6541
1.2187 | -89.4638
1.2225 | -89.1489
1.2262 | -88.8043
1.2299 | -88.6703
1.2336 | -88.8042
1.2374 | -89.0421
1.2411 | -89.1822
1.2448 | -89.2187
1.2485 | -89.2923
1.2523 | -89.5451
1.256 | -89.9737
1.2597 | -90.2197
1.2635 | -89.9105
1.2672 | -89.2941
1.2709 | -88.8467
1.2746 | -88.7657
1.2784 | -89.017
1.2821 | -89.4825
1.2858 | -89.9964
1.2895 | -90.3518
1.2933 | -90.3423
1.297 | -89.9609
1.3007 | -89.5489
1.3044 | -89.3492
1.3082 | -89.2728
1.3119 | -89.239
1.3156 | -89.2122
1.3194 | -89.1596
1.3231 | -89.2687
1.3268 | -89.6636
1.3305 | -90.0929
1.3343 | -90.334
1.338 | -90.5271
1.3417 | -90.6078
1.3454 | -90.2539
1.3492 | -89.6731
1.3529 | -89.3982
1.3566 | -89.631
1.3604 | -90.1613
1.3641 | -90.6303
1.3678 | -90.8858
1.3715 | -91.0297
1.3753 | -91.1689
1.379 | -91.26
1.3827 | -91.1914
1.3864 | -90.9813
1.3902 | -90.8711
1.3939 | -91.0204
1.3976 | -91.2502
1.4014 | -91.251
1.4051 | -90.9239
1.4088 | -90.3911
1.4125 | -89.899
1.4163 | -89.5957
1.42 | -89.4869
1.4237 | -89.5463
1.4274 | -89.7615
1.4312 | -90.0059
1.4349 | -90.1651
1.4386 | -90.3164
1.4423 | -90.4636
1.4461 | -90.4048
1.4498 | -90.203
1.4535 | -90.255
1.4573 | -90.6327
1.461 | -90.8406
1.4647 | -90.6104
1.4684 | -90.3766
1.4722 | -90.4994
1.4759 | -90.8258
1.4796 | -90.9885
1.4833 | -91.0056
1.4871 | -91.0083
1.4908 | -90.8526
1.4945 | -90.5419
1.4983 | -90.3122
1.502 | -90.2921
1.5057 | -90.41
1.5094 | -90.4272
1.5132 | -90.3155
1.5169 | -90.3873
1.5206 | -90.7946
1.5243 | -91.3458
1.5281 | -91.7092
1.5318 | -91.7307
1.5355 | -91.5717
1.5393 | -91.3812
1.543 | -91.1804
1.5467 | -91.1071
1.5504 | -91.2732
1.5542 | -91.4673
1.5579 | -91.3988
1.5616 | -91.0947
1.5653 | -90.798
1.5691 | -90.686
1.5728 | -90.8236
1.5765 | -91.2005
1.5802 | -91.5552
1.584 | -91.5509
1.5877 | -91.3332
1.5914 | -91.1666
1.5952 | -91.0554
1.5989 | -90.9135
1.6026 | -90.7549
1.6063 | -90.7786
1.6101 | -91.0815
1.6138 | -91.3117
1.6175 | -91.4713
1.6212 | -91.5089
1.625 | -91.4581
1.6287 | -91.4242
1.6324 | -91.4884
1.6362 | -91.5968
1.6399 | -91.646
1.6436 | -91.5937
1.6473 | -91.4706
1.6511 | -91.3466
1.6548 | -91.314
1.6585 | -91.4345
1.6622 | -91.7185
1.666 | -92.1072
1.6697 | -92.4384
1.6734 | -92.5329
1.6771 | -92.3809
1.6809 | -92.1723
1.6846 | -92.0516
1.6883 | -92.0476
1.6921 | -92.1496
1.6958 | -92.2906
1.6995 | -92.3976
1.7032 | -92.4393
1.707 | -92.4292
1.7107 | -92.366
1.7144 | -92.2233
1.7181 | -92.0762
1.7219 | -92.0368
1.7256 | -92.1395
1.7293 | -92.3674
1.7331 | -92.6414
1.7368 | -92.8088
1.7405 | -92.7446
1.7442 | -92.48
1.748 | -92.1726
1.7517 | -91.9436
1.7554 | -91.8606
1.7591 | -91.9374
1.7629 | -92.0951
1.7666 | -92.2121
1.7703 | -92.2408
1.7741 | -92.2173
1.7778 | -92.2193
1.7815 | -92.2854
1.7852 | -92.4337
1.789 | -92.6742
1.7927 | -93.0055
1.7964 | -93.3797
1.8001 | -93.7341
1.8039 | -93.9238
1.8076 | -93.8435
1.8113 | -93.5615
1.815 | -93.2818
1.8188 | -93.1458
1.8225 | -93.1741
1.8262 | -93.2722
1.83 | -93.3252
1.8337 | -93.2525
1.8374 | -93.0714
1.8411 | -92.8661
1.8449 | -92.746
1.8486 | -92.7515
1.8523 | -92.8117
1.856 | -92.7922
1.8598 | -92.7054
1.8635 | -92.6876
1.8672 | -92.8519
1.871 | -93.1594
1.8747 | -93.4382
1.8784 | -93.5747
1.8821 | -93.6148
1.8859 | -93.6057
1.8896 | -93.5408
1.8933 | -93.4305
1.897 | -93.3412
1.9008 | -93.3249
1.9045 | -93.3968
1.9082 | -93.4917
1.912 | -93.4915
1.9157 | -93.3248
1.9194 | -93.1065
1.9231 | -93.0115
1.9269 | -93.1132
1.9306 | -93.3688
1.9343 | -93.6434
1.938 | -93.8062
1.9418 | -93.8225
1.9455 | -93.7026
1.9492 | -93.4755
1.9529 | -93.2034
1.9567 | -92.9859
1.9604 | -92.8839
1.9641 | -92.8688
1.9679 | -92.8435
1.9716 | -92.739
1.9753 | -92.615
1.979 | -92.5846
1.9828 | -92.6302
1.9865 | -92.6834
1.9902 | -92.7777
1.9939 | -92.9872
1.9977 | -93.2965
2.0014 | -93.5634
2.0051 | -93.6523
2.0089 | -93.5354
2.0126 | -93.3333
2.0163 | -93.1521
2.02 | -92.9965
2.0238 | -92.8344
2.0275 | -92.7084
2.0312 | -92.6761
2.0349 | -92.7361
2.0387 | -92.8296
2.0424 | -92.9535
2.0461 | -93.1177
2.0498 | -93.3211
2.0536 | -93.4905
2.0573 | -93.5774
2.061 | -93.5624
2.0648 | -93.4897
2.0685 | -93.4263
2.0722 | -93.4194
2.0759 | -93.4544
2.0797 | -93.4999
2.0834 | -93.5281
2.0871 | -93.5834
2.0908 | -93.7038
2.0946 | -93.8261
2.0983 | -93.8501
2.102 | -93.764
2.1058 | -93.6362
2.1095 | -93.5846
2.1132 | -93.7003
2.1169 | -93.9586
2.1207 | -94.2141
2.1244 | -94.3527
2.1281 | -94.4264
2.1318 | -94.4892
2.1356 | -94.5499
2.1393 | -94.6005
2.143 | -94.6509
2.1468 | -94.7151
2.1505 | -94.7306
2.1542 | -94.6879
2.1579 | -94.5592
2.1617 | -94.3584
2.1654 | -94.1217
2.1691 | -93.8823
2.1728 | -93.6564
2.1766 | -93.466
2.1803 | -93.3245
2.184 | -93.2452
2.1877 | -93.2432
2.1915 | -93.3305
2.1952 | -93.4943
2.1989 | -93.668
2.2027 | -93.7866
2.2064 | -93.8202
2.2101 | -93.8031
2.2138 | -93.7704
2.2176 | -93.7489
2.2213 | -93.7523
2.225 | -93.773
2.2287 | -93.7872
2.2325 | -93.7919
2.2362 | -93.8008
2.2399 | -93.8426
2.2437 | -93.9071
2.2474 | -93.9674
2.2511 | -93.9915
2.2548 | -93.9902
2.2586 | -93.9856
2.2623 | -93.9948
2.266 | -94.018
2.2697 | -94.0604
2.2735 | -94.1038
2.2772 | -94.1319
2.2809 | -94.1367
2.2846 | -94.1536
2.2884 | -94.2204
2.2921 | -94.3562
2.2958 | -94.5145
2.2996 | -94.6231
2.3033 | -94.6162
2.307 | -94.5286
2.3107 | -94.4388
2.3145 | -94.4256
2.3182 | -94.5119
2.3219 | -94.6739
2.3256 | -94.8313
2.3294 | -94.9327
2.3331 | -94.9476
2.3368 | -94.8717
2.3406 | -94.738
2.3443 | -94.5958
2.348 | -94.4788
2.3517 | -94.4208
2.3555 | -94.4148
2.3592 | -94.4288
2.3629 | -94.4391
2.3666 | -94.425
2.3704 | -94.3563
2.3741 | -94.2527
2.3778 | -94.172
2.3816 | -94.1555
2.3853 | -94.2272
2.389 | -94.3682
2.3927 | -94.5381
2.3965 | -94.7162
2.4002 | -94.8868
2.4039 | -95.0396
2.4076 | -95.1729
2.4114 | -95.2605
2.4151 | -95.2891
2.4188 | -95.2871
2.4225 | -95.278
2.4263 | -95.236
2.43 | -95.1377
2.4337 | -94.9871
2.4375 | -94.8131
2.4412 | -94.6794
2.4449 | -94.6339
2.4486 | -94.6827
2.4524 | -94.7892
2.4561 | -94.9047
2.4598 | -94.988
2.4635 | -95.0425
2.4673 | -95.0658
2.471 | -95.0385
2.4747 | -94.9779
2.4785 | -94.9253
2.4822 | -94.9067
2.4859 | -94.9212
2.4896 | -94.9577
2.4934 | -95.0143
2.4971 | -95.086
2.5008 | -95.1689
2.5045 | -95.2437
2.5083 | -95.2839
2.512 | -95.2788
2.5157 | -95.2365
2.5194 | -95.1633
2.5232 | -95.0765
2.5269 | -95.0235
2.5306 | -95.0644
2.5344 | -95.2398
2.5381 | -95.4983
2.5418 | -95.7488
2.5455 | -95.877
2.5493 | -95.8457
2.553 | -95.6993
2.5567 | -95.5191
2.5604 | -95.3935
2.5642 | -95.371
2.5679 | -95.4397
2.5716 | -95.5438
2.5754 | -95.6141
2.5791 | -95.6381
2.5828 | -95.6127
2.5865 | -95.5501
2.5903 | -95.4483
2.594 | -95.3019
2.5977 | -95.1328
2.6014 | -94.9925
2.6052 | -94.9224
2.6089 | -94.9413
2.6126 | -95.0412
2.6164 | -95.1725
2.6201 | -95.3131
2.6238 | -95.4627
2.6275 | -95.6149
2.6313 | -95.7291
2.635 | -95.778
2.6387 | -95.7682
2.6424 | -95.7282
2.6462 | -95.6742
2.6499 | -95.5942
2.6536 | -95.5011
2.6573 | -95.4451
2.6611 | -95.4277
2.6648 | -95.439
2.6685 | -95.4423
2.6723 | -95.4336
2.676 | -95.4431
2.6797 | -95.4825
2.6834 | -95.5363
2.6872 | -95.59
2.6909 | -95.6301
2.6946 | -95.6716
2.6983 | -95.6993
2.7021 | -95.7157
2.7058 | -95.7214
2.7095 | -95.7178
2.7133 | -95.7194
2.717 | -95.7266
2.7207 | -95.74
2.7244 | -95.7656
2.7282 | -95.7977
2.7319 | -95.8246
2.7356 | -95.835
2.7393 | -95.8009
2.7431 | -95.7075
2.7468 | -95.5713
2.7505 | -95.4178
2.7543 | -95.2975
2.758 | -95.2446
2.7617 | -95.277
2.7654 | -95.3788
2.7692 | -95.5152
2.7729 | -95.6504
2.7766 | -95.7654
2.7803 | -95.8693
2.7841 | -95.9722
2.7878 | -96.0852
2.7915 | -96.2131
2.7952 | -96.3399
2.799 | -96.4481
2.8027 | -96.5293
2.8064 | -96.5757
2.8102 | -96.5802
2.8139 | -96.5517
2.8176 | -96.4971
2.8213 | -96.4281
2.8251 | -96.3441
2.8288 | -96.252
2.8325 | -96.1739
2.8362 | -96.1096
2.84 | -96.0611
2.8437 | -96.024
2.8474 | -95.9885
2.8512 | -95.9498
2.8549 | -95.9197
2.8586 | -95.9068
2.8623 | -95.9126
2.8661 | -95.9672
2.8698 | -96.0862
2.8735 | -96.2836
2.8772 | -96.5428
2.881 | -96.7971
2.8847 | -96.9724
2.8884 | -97.0232
2.8921 | -96.9637
2.8959 | -96.8242
2.8996 | -96.671
2.9033 | -96.5622
2.9071 | -96.5186
2.9108 | -96.5387
2.9145 | -96.5968
2.9182 | -96.6682
2.922 | -96.718
2.9257 | -96.7312
2.9294 | -96.7176
2.9331 | -96.6947
2.9369 | -96.6451
2.9406 | -96.5719
2.9443 | -96.4824
2.9481 | -96.377
2.9518 | -96.2911
2.9555 | -96.2507
2.9592 | -96.2562
2.963 | -96.3016
2.9667 | -96.3756
2.9704 | -96.4654
2.9741 | -96.568
2.9779 | -96.6757
2.9816 | -96.7826
2.9853 | -96.8988
2.9891 | -97.0167
2.9928 | -97.1211
2.9965 | -97.175
3.0002 | -97.1343
3.004 | -97.0098
3.0077 | -96.8447
3.0114 | -96.6623
3.0151 | -96.5
3.0189 | -96.3722
3.0226 | -96.2817
3.0263 | -96.2361
3.03 | -96.2322
3.0338 | -96.2657
3.0375 | -96.316
3.0412 | -96.3607
3.045 | -96.3823
3.0487 | -96.3897
3.0524 | -96.3831
3.0561 | -96.392
3.0599 | -96.4506
3.0636 | -96.5474
3.0673 | -96.6631
3.071 | -96.7492
3.0748 | -96.8029
3.0785 | -96.8259
3.0822 | -96.8269
3.086 | -96.8212
3.0897 | -96.8195
3.0934 | -96.8178
3.0971 | -96.8159
3.1009 | -96.8174
3.1046 | -96.8081
3.1083 | -96.8064
3.112 | -96.8078
3.1158 | -96.8219
3.1195 | -96.8756
3.1232 | -96.9658
3.127 | -97.1102
3.1307 | -97.2932
3.1344 | -97.4747
3.1381 | -97.6115
3.1419 | -97.6555
3.1456 | -97.5887
3.1493 | -97.4392
3.153 | -97.2514
3.1568 | -97.0696
3.1605 | -96.9473
3.1642 | -96.8699
3.1679 | -96.8339
3.1717 | -96.845
3.1754 | -96.8837
3.1791 | -96.9553
3.1829 | -97.0487
3.1866 | -97.1596
3.1903 | -97.2977
3.194 | -97.4512
3.1978 | -97.5996
3.2015 | -97.7141
3.2052 | -97.7673
3.2089 | -97.7686
3.2127 | -97.7507
3.2164 | -97.7238
3.2201 | -97.6948
3.2239 | -97.7172
3.2276 | -97.7268
3.2313 | -97.7448
3.235 | -97.7649
3.2388 | -97.7663
3.2425 | -97.743
3.2462 | -97.6985
3.2499 | -97.6331
3.2537 | -97.5663
3.2574 | -97.518
3.2611 | -97.4821
3.2649 | -97.4699
3.2686 | -97.4807
3.2723 | -97.5131
3.276 | -97.5635
3.2798 | -97.635
3.2835 | -97.7194
3.2872 | -97.8109
3.2909 | -97.908
3.2947 | -98.0088
3.2984 | -98.0955
3.3021 | -98.1552
3.3058 | -98.1827
3.3096 | -98.1603
3.3133 | -98.0943
3.317 | -98.0238
3.3208 | -97.9654
3.3245 | -97.939
3.3282 | -97.9537
3.3319 | -97.9857
3.3357 | -98.0099
3.3394 | -98.0154
3.3431 | -97.9919
3.3468 | -97.9405
3.3506 | -97.8607
3.3543 | -97.7625
3.358 | -97.6739
3.3618 | -97.6027
3.3655 | -97.5497
3.3692 | -97.5153
3.3729 | -97.4957
3.3767 | -97.4889
3.3804 | -97.4951
3.3841 | -97.5122
3.3878 | -97.558
3.3916 | -97.6261
3.3953 | -97.6988
3.399 | -97.7734
3.4027 | -97.8219
3.4065 | -97.8375
3.4102 | -97.8221
3.4139 | -97.7478
3.4177 | -97.6314
3.4214 | -97.5022
3.4251 | -97.3678
3.4288 | -97.2492
3.4326 | -97.1615
3.4363 | -97.1135
3.44 | -97.0962
3.4437 | -97.1103
3.4475 | -97.1597
3.4512 | -97.2234
3.4549 | -97.2809
3.4587 | -97.3328
3.4624 | -97.3773
3.4661 | -97.4214
3.4698 | -97.5004
3.4736 | -97.6099
3.4773 | -97.7231
3.481 | -97.8391
3.4847 | -97.9278
3.4885 | -97.9873
3.4922 | -98.014
3.4959 | -98.0089
3.4997 | -97.9605
3.5034 | -97.8781
3.5071 | -97.7917
3.5108 | -97.7134
3.5146 | -97.651
3.5183 | -97.6006
3.522 | -97.5691
3.5257 | -97.5461
3.5295 | -97.5274
3.5332 | -97.5051
3.5369 | -97.4775
3.5406 | -97.4512
3.5444 | -97.4456
3.5481 | -97.4708
3.5518 | -97.5252
3.5556 | -97.596
3.5593 | -97.6614
3.563 | -97.7039
3.5667 | -97.7118
3.5705 | -97.6921
3.5742 | -97.6733
3.5779 | -97.6775
3.5816 | -97.7166
3.5854 | -97.7905
3.5891 | -97.8868
3.5928 | -97.9707
3.5966 | -98.0335
3.6003 | -98.0572
3.604 | -98.054
3.6077 | -98.0564
3.6115 | -98.0841
3.6152 | -98.1426
3.6189 | -98.2283
3.6226 | -98.3401
3.6264 | -98.4511
3.6301 | -98.5379
3.6338 | -98.5884
3.6375 | -98.6241
3.6413 | -98.6509
3.645 | -98.6566
3.6487 | -98.6386
3.6525 | -98.5866
3.6562 | -98.5035
3.6599 | -98.4152
3.6636 | -98.3439
3.6674 | -98.2989
3.6711 | -98.2941
3.6748 | -98.3204
3.6785 | -98.3328
3.6823 | -98.3202
3.686 | -98.2849
3.6897 | -98.2346
3.6935 | -98.172
3.6972 | -98.1022
3.7009 | -98.0321
3.7046 | -97.9552
3.7084 | -97.88
3.7121 | -97.8206
3.7158 | -97.7807
3.7195 | -97.7617
3.7233 | -97.7649
3.727 | -97.7934
3.7307 | -97.8483
3.7345 | -97.9449
3.7382 | -98.0699
3.7419 | -98.2012
3.7456 | -98.3283
3.7494 | -98.4407
3.7531 | -98.5351
3.7568 | -98.6109
3.7605 | -98.6481
3.7643 | -98.6902
3.768 | -98.7174
3.7717 | -98.7255
3.7754 | -98.7156
3.7792 | -98.6996
3.7829 | -98.6778
3.7866 | -98.6511
3.7904 | -98.6146
3.7941 | -98.5684
3.7978 | -98.528
3.8015 | -98.4936
3.8053 | -98.4693
3.809 | -98.4639
3.8127 | -98.4712
3.8164 | -98.4904
3.8202 | -98.5199
3.8239 | -98.5624
3.8276 | -98.6133
3.8314 | -98.6626
3.8351 | -98.705
3.8388 | -98.7528
3.8425 | -98.7818
3.8463 | -98.8088
3.85 | -98.8132
3.8537 | -98.8102
3.8574 | -98.7996
3.8612 | -98.7663
3.8649 | -98.723
3.8686 | -98.6825
3.8723 | -98.6457
3.8761 | -98.6296
3.8798 | -98.6255
3.8835 | -98.6242
3.8873 | -98.6305
3.891 | -98.6506
3.8947 | -98.6778
3.8984 | -98.7071
3.9022 | -98.7194
3.9059 | -98.7259
3.9096 | -98.7277
3.9133 | -98.7326
3.9171 | -98.7341
3.9208 | -98.7346
3.9245 | -98.7453
3.9283 | -98.7573
3.932 | -98.7557
3.9357 | -98.7481
3.9394 | -98.7361
3.9432 | -98.7311
3.9469 | -98.7345
3.9506 | -98.7355
3.9543 | -98.7324
3.9581 | -98.7289
3.9618 | -98.7284
3.9655 | -98.7327
3.9693 | -98.7449
3.973 | -98.7516
3.9767 | -98.762
3.9804 | -98.7812
3.9842 | -98.7878
3.9879 | -98.7858
3.9916 | -98.7837
3.9953 | -98.7833
3.9991 | -98.791
4.0028 | -98.8084
4.0065 | -98.8437
4.0102 | -98.9057
4.014 | -98.9963
4.0177 | -99.0946
4.0214 | -99.183
4.0252 | -99.2459
4.0289 | -99.2844
4.0326 | -99.2984
4.0363 | -99.2872
4.0401 | -99.2718
4.0438 | -99.2497
4.0475 | -99.2361
4.0512 | -99.2242
4.055 | -99.2292
4.0587 | -99.254
4.0624 | -99.2951
4.0662 | -99.3611
4.0699 | -99.4269
4.0736 | -99.4904
4.0773 | -99.5342
4.0811 | -99.5454
4.0848 | -99.5158
4.0885 | -99.4593
4.0922 | -99.3824
4.096 | -99.3106
4.0997 | -99.2755
4.1034 | -99.2724
4.1072 | -99.2956
4.1109 | -99.3452
4.1146 | -99.4042
4.1183 | -99.4531
4.1221 | -99.4815
4.1258 | -99.48
4.1295 | -99.4449
4.1332 | -99.4046
4.137 | -99.3566
4.1407 | -99.3169
4.1444 | -99.2946
4.1481 | -99.2733
4.1519 | -99.2579
4.1556 | -99.2476
4.1593 | -99.246
4.1631 | -99.2539
4.1668 | -99.2694
4.1705 | -99.3081
4.1742 | -99.3464
4.178 | -99.3748
4.1817 | -99.3856
4.1854 | -99.3792
4.1891 | -99.3595
4.1929 | -99.3352
4.1966 | -99.3131
4.2003 | -99.3088
4.2041 | -99.3337
4.2078 | -99.3865
4.2115 | -99.47
4.2152 | -99.5647
4.219 | -99.6727
4.2227 | -99.7722
4.2264 | -99.8529
4.2301 | -99.9144
4.2339 | -99.934
4.2376 | -99.93
4.2413 | -99.9083
4.245 | -99.8754
4.2488 | -99.8363
4.2525 | -99.7961
4.2562 | -99.758
4.26 | -99.7256
4.2637 | -99.6992
4.2674 | -99.6792
4.2711 | -99.665
4.2749 | -99.6462
4.2786 | -99.6288
4.2823 | -99.6029
4.286 | -99.5682
4.2898 | -99.5245
4.2935 | -99.4645
4.2972 | -99.3995
4.301 | -99.3442
4.3047 | -99.3027
4.3084 | -99.2795
4.3121 | -99.2751
4.3159 | -99.2842
4.3196 | -99.3048
4.3233 | -99.3325
4.327 | -99.3584
4.3308 | -99.3804
4.3345 | -99.3959
4.3382 | -99.4033
4.342 | -99.4043
4.3457 | -99.4053
4.3494 | -99.4082
4.3531 | -99.4102
4.3569 | -99.4141
4.3606 | -99.4352
4.3643 | -99.4641
4.368 | -99.5046
4.3718 | -99.5478
4.3755 | -99.5992
4.3792 | -99.6607
4.3829 | -99.7369
4.3867 | -99.8257
4.3904 | -99.9228
4.3941 | -100.0266
4.3979 | -100.1317
4.4016 | -100.2343
4.4053 | -100.3181
4.409 | -100.3707
4.4128 | -100.3988
4.4165 | -100.3886
4.4202 | -100.3539
4.4239 | -100.3007
4.4277 | -100.2265
4.4314 | -100.1357
4.4351 | -100.0527
4.4389 | -99.9843
4.4426 | -99.9361
4.4463 | -99.9131
4.45 | -99.9123
4.4538 | -99.9341
4.4575 | -99.9676
4.4612 | -99.9915
4.4649 | -100.0025
4.4687 | -99.9928
4.4724 | -99.9527
4.4761 | -99.8985
4.4799 | -99.8428
4.4836 | -99.7871
4.4873 | -99.7476
4.491 | -99.7374
4.4948 | -99.7635
4.4985 | -99.826
4.5022 | -99.9165
4.5059 | -100.0215
4.5097 | -100.1335
4.5134 | -100.2399
4.5171 | -100.3302
4.5208 | -100.411
4.5246 | -100.4738
4.5283 | -100.5161
4.532 | -100.5458
4.5358 | -100.5576
4.5395 | -100.5632
4.5432 | -100.5574
4.5469 | -100.5329
4.5507 | -100.4917
4.5544 | -100.4288
4.5581 | -100.3586
4.5618 | -100.2883
4.5656 | -100.2257
4.5693 | -100.183
4.573 | -100.1625
4.5768 | -100.1646
4.5805 | -100.1811
4.5842 | -100.2194
4.5879 | -100.2717
4.5917 | -100.3266
4.5954 | -100.4017
4.5991 | -100.5003
4.6028 | -100.6097
4.6066 | -100.7145
4.6103 | -100.8095
4.614 | -100.8904
4.6178 | -100.9491
4.6215 | -100.9715
4.6252 | -100.9712
4.6289 | -100.9348
4.6327 | -100.8618
4.6364 | -100.7698
4.6401 | -100.6617
4.6438 | -100.5575
4.6476 | -100.4619
4.6513 | -100.3738
4.655 | -100.3053
4.6587 | -100.2697
4.6625 | -100.2555
4.6662 | -100.2511
4.6699 | -100.25
4.6737 | -100.2458
4.6774 | -100.2387
4.6811 | -100.2129
4.6848 | -100.1738
4.6886 | -100.1281
4.6923 | -100.0791
4.696 | -100.0291
4.6997 | -100.0003
4.7035 | -99.9951
4.7072 | -100.0056
4.7109 | -100.0401
4.7147 | -100.0864
4.7184 | -100.1392
4.7221 | -100.1866
4.7258 | -100.2144
4.7296 | -100.2213
4.7333 | -100.2246
4.737 | -100.2283
4.7407 | -100.2384
4.7445 | -100.2603
4.7482 | -100.2913
4.7519 | -100.341
4.7556 | -100.3988
4.7594 | -100.4519
4.7631 | -100.5159
4.7668 | -100.573
4.7706 | -100.6038
4.7743 | -100.6143
4.778 | -100.6098
4.7817 | -100.5871
4.7855 | -100.5567
4.7892 | -100.5282
4.7929 | -100.5247
4.7966 | -100.5322
4.8004 | -100.5398
4.8041 | -100.5433
4.8078 | -100.5291
4.8116 | -100.5042
4.8153 | -100.4666
4.819 | -100.4271
4.8227 | -100.4017
4.8265 | -100.4097
4.8302 | -100.4594
4.8339 | -100.5398
4.8376 | -100.6674
4.8414 | -100.7543
4.8451 | -100.847
4.8488 | -100.9229
4.8526 | -100.9589
4.8563 | -100.9579
4.86 | -100.9209
4.8637 | -100.8477
4.8675 | -100.7453
4.8712 | -100.633
4.8749 | -100.5307
4.8786 | -100.4408
4.8824 | -100.3768
4.8861 | -100.3378
4.8898 | -100.3339
4.8935 | -100.3595
4.8973 | -100.4094
4.901 | -100.4852
4.9047 | -100.5809
4.9085 | -100.6859
4.9122 | -100.8012
4.9159 | -100.9163
4.9196 | -101.0075
4.9234 | -101.0697
4.9271 | -101.1088
4.9308 | -101.1315
4.9345 | -101.121
4.9383 | -101.0831
4.942 | -101.0192
4.9457 | -100.9284
4.9495 | -100.8204
4.9532 | -100.7098
4.9569 | -100.5993
4.9606 | -100.4991
4.9644 | -100.4201
4.9681 | -100.3596
4.9718 | -100.316
4.9755 | -100.2909
4.9793 | -100.2792
4.983 | -100.2875
4.9867 | -100.3182
4.9905 | -100.3709
4.9942 | -100.4481
4.9979 | -100.5454
5.0016 | -100.6555
5.0054 | -100.7667
5.0091 | -100.8679
5.0128 | -100.9453
5.0165 | -100.988
5.0203 | -101.0007
5.024 | -100.9783
5.0277 | -100.9253
5.0314 | -100.8457
5.0352 | -100.76
5.0389 | -100.6672
5.0426 | -100.5808
5.0464 | -100.5145
5.0501 | -100.4851
5.0538 | -100.4904
5.0575 | -100.5341
5.0613 | -100.6028
5.065 | -100.7015
5.0687 | -100.8093
5.0724 | -100.9162
5.0762 | -101.0169
5.0799 | -101.0931
5.0836 | -101.1444
5.0874 | -101.1674
5.0911 | -101.1707
5.0948 | -101.158
5.0985 | -101.1248
5.1023 | -101.0871
5.106 | -101.0592
5.1097 | -101.04
5.1134 | -101.0309
5.1172 | -101.037
5.1209 | -101.0454
5.1246 | -101.0425
5.1283 | -101.0252
5.1321 | -100.9933
5.1358 | -100.9437
5.1395 | -100.8919
5.1433 | -100.8326
5.147 | -100.7872
5.1507 | -100.7597
5.1544 | -100.7448
5.1582 | -100.7521
5.1619 | -100.7745
5.1656 | -100.812
5.1693 | -100.8607
5.1731 | -100.9163
5.1768 | -100.9803
5.1805 | -101.0475
5.1843 | -101.1085
5.188 | -101.1665
5.1917 | -101.2285
5.1954 | -101.2875
5.1992 | -101.3357
5.2029 | -101.38
5.2066 | -101.4209
5.2103 | -101.4566
5.2141 | -101.4845
5.2178 | -101.5005
5.2215 | -101.5039
5.2252 | -101.4979
5.229 | -101.4769
5.2327 | -101.4455
5.2364 | -101.3997
5.2402 | -101.351
5.2439 | -101.3097
5.2476 | -101.2737
5.2513 | -101.2446
5.2551 | -101.2247
5.2588 | -101.2131
5.2625 | -101.2066
5.2662 | -101.2078
5.27 | -101.206
5.2737 | -101.2097
5.2774 | -101.2203
5.2812 | -101.2308
5.2849 | -101.2347
5.2886 | -101.246
5.2923 | -101.2613
5.2961 | -101.2696
5.2998 | -101.282
5.3035 | -101.2845
5.3072 | -101.2793
5.311 | -101.2718
5.3147 | -101.2611
5.3184 | -101.2313
5.3222 | -101.1897
5.3259 | -101.1467
5.3296 | -101.1086
5.3333 | -101.0826
5.3371 | -101.0592
5.3408 | -101.0411
5.3445 | -101.0308
5.3482 | -101.0328
5.352 | -101.0337
5.3557 | -101.0374
5.3594 | -101.0515
5.3631 | -101.0665
5.3669 | -101.0706
5.3706 | -101.0747
5.3743 | -101.0801
5.3781 | -101.0974
5.3818 | -101.1145
5.3855 | -101.1401
5.3892 | -101.1701
5.393 | -101.2005
5.3967 | -101.2342
5.4004 | -101.2778
5.4041 | -101.314
5.4079 | -101.3476
5.4116 | -101.3676
5.4153 | -101.371
5.4191 | -101.3692
5.4228 | -101.354
5.4265 | -101.3223
5.4302 | -101.2822
5.434 | -101.2421
5.4377 | -101.1958
5.4414 | -101.1469
5.4451 | -101.1091
5.4489 | -101.0863
5.4526 | -101.0737
5.4563 | -101.0743
5.4601 | -101.0871
5.4638 | -101.1115
5.4675 | -101.1487
5.4712 | -101.1982
5.475 | -101.2522
5.4787 | -101.304
5.4824 | -101.3477
5.4861 | -101.3811
5.4899 | -101.4117
5.4936 | -101.4363
5.4973 | -101.4617
5.501 | -101.4912
5.5048 | -101.5219
5.5085 | -101.5585
5.5122 | -101.6001
5.516 | -101.6461
5.5197 | -101.6996
5.5234 | -101.7589
5.5271 | -101.8142
5.5309 | -101.865
5.5346 | -101.9119
5.5383 | -101.9588
5.542 | -101.9965
5.5458 | -102.0236
5.5495 | -102.046
5.5532 | -102.0536
5.557 | -102.0557
5.5607 | -102.059
5.5644 | -102.0594
5.5681 | -102.0594
5.5719 | -102.0594
5.5756 | -102.0624
5.5793 | -102.0619
5.583 | -102.0625
5.5868 | -102.0603
5.5905 | -102.054
5.5942 | -102.05
5.5979 | -102.0473
5.6017 | -102.0523
5.6054 | -102.0603
5.6091 | -102.0744
5.6129 | -102.0951
5.6166 | -102.1166
5.6203 | -102.1348
5.624 | -102.1406
5.6278 | -102.1487
5.6315 | -102.1567
5.6352 | -102.1572
5.6389 | -102.1648
5.6427 | -102.1689
5.6464 | -102.1604
5.6501 | -102.1425
5.6539 | -102.1201
5.6576 | -102.0914
5.6613 | -102.0557
5.665 | -102.0194
5.6688 | -101.9823
5.6725 | -101.9481
5.6762 | -101.9231
5.6799 | -101.9037
5.6837 | -101.8886
5.6874 | -101.8767
5.6911 | -101.8731
5.6949 | -101.866
5.6986 | -101.8626
5.7023 | -101.8646
5.706 | -101.8651
5.7098 | -101.8559
5.7135 | -101.8366
5.7172 | -101.8115
5.7209 | -101.775
5.7247 | -101.7258
5.7284 | -101.6687
5.7321 | -101.6103
5.7358 | -101.5504
5.7396 | -101.4956
5.7433 | -101.451
5.747 | -101.4236
5.7508 | -101.4104
5.7545 | -101.4109
5.7582 | -101.4307
5.7619 | -101.4618
5.7657 | -101.4977
5.7694 | -101.5391
5.7731 | -101.5886
5.7768 | -101.6378
5.7806 | -101.6914
5.7843 | -101.74
5.788 | -101.7899
5.7918 | -101.8486
5.7955 | -101.921
5.7992 | -101.9888
5.8029 | -102.0577
5.8067 | -102.1306
5.8104 | -102.1927
5.8141 | -102.2419
5.8178 | -102.2794
5.8216 | -102.3025
5.8253 | -102.3055
5.829 | -102.2949
5.8328 | -102.2758
5.8365 | -102.2599
5.8402 | -102.2352
5.8439 | -102.2039
5.8477 | -102.1749
5.8514 | -102.1477
5.8551 | -102.1185
5.8588 | -102.0878
5.8626 | -102.0597
5.8663 | -102.0341
5.87 | -102.016
5.8737 | -101.9927
5.8775 | -101.9715
5.8812 | -101.9427
5.8849 | -101.9142
5.8887 | -101.8745
5.8924 | -101.8231
5.8961 | -101.7686
5.8998 | -101.7099
5.9036 | -101.6519
5.9073 | -101.6
5.911 | -101.5908
5.9147 | -101.5799
5.9185 | -101.5897
5.9222 | -101.6193
5.9259 | -101.668
5.9297 | -101.7426
5.9334 | -101.8316
5.9371 | -101.9286
5.9408 | -102.0358
5.9446 | -102.1472
5.9483 | -102.2525
5.952 | -102.3469
5.9557 | -102.4164
5.9595 | -102.462
5.9632 | -102.4914
5.9669 | -102.4949
5.9707 | -102.4849
5.9744 | -102.4565
5.9781 | -102.4144
5.9818 | -102.3752
5.9856 | -102.3361
5.9893 | -102.2992
5.993 | -102.2682
5.9967 | -102.2464
6.0005 | -102.2363
6.0042 | -102.2271
6.0079 | -102.2184
6.0116 | -102.213
6.0154 | -102.2048
6.0191 | -102.1992
6.0228 | -102.1934
6.0266 | -102.1895
6.0303 | -102.1864
6.034 | -102.1803
6.0377 | -102.1698
6.0415 | -102.1628
6.0452 | -102.1643
6.0489 | -102.1742
6.0526 | -102.1776
6.0564 | -102.1838
6.0601 | -102.1945
6.0638 | -102.2047
6.0676 | -102.2176
6.0713 | -102.2306
6.075 | -102.2368
6.0787 | -102.2385
6.0825 | -102.2338
6.0862 | -102.2281
6.0899 | -102.2191
6.0936 | -102.2067
6.0974 | -102.1963
6.1011 | -102.1874
6.1048 | -102.1835
6.1085 | -102.1872
6.1123 | -102.1903
6.116 | -102.1981
6.1197 | -102.213
6.1235 | -102.2374
6.1272 | -102.2734
6.1309 | -102.316
6.1346 | -102.3636
6.1384 | -102.4142
6.1421 | -102.4667
6.1458 | -102.5285
6.1495 | -102.5879
6.1533 | -102.6369
6.157 | -102.679
6.1607 | -102.7099
6.1645 | -102.7339
6.1682 | -102.7491
6.1719 | -102.7539
6.1756 | -102.759
6.1794 | -102.7578
6.1831 | -102.755
6.1868 | -102.752
6.1905 | -102.7489
6.1943 | -102.745
6.198 | -102.7395
6.2017 | -102.7387
6.2055 | -102.7385
6.2092 | -102.7403
6.2129 | -102.7495
6.2166 | -102.7579
6.2204 | -102.7597
6.2241 | -102.7708
6.2278 | -102.7864
6.2315 | -102.8015
6.2353 | -102.8159
6.239 | -102.8231
6.2427 | -102.825
6.2464 | -102.8274
6.2502 | -102.8189
6.2539 | -102.8015
6.2576 | -102.7754
6.2614 | -102.7393
6.2651 | -102.6993
6.2688 | -102.6501
6.2725 | -102.5945
6.2763 | -102.5371
6.28 | -102.4798
6.2837 | -102.4156
6.2874 | -102.3522
6.2912 | -102.2879
6.2949 | -102.2243
6.2986 | -102.1704
6.3024 | -102.115
6.3061 | -102.0658
6.3098 | -102.0341
6.3135 | -102.0221
6.3173 | -102.0223
6.321 | -102.0301
6.3247 | -102.0528
6.3284 | -102.0861
6.3322 | -102.1288
6.3359 | -102.1784
6.3396 | -102.2228
6.3434 | -102.2684
6.3471 | -102.305
6.3508 | -102.3357
6.3545 | -102.3608
6.3583 | -102.377
6.362 | -102.3852
6.3657 | -102.3898
6.3694 | -102.3915
6.3732 | -102.3869
6.3769 | -102.3887
6.3806 | -102.3854
6.3843 | -102.3835
6.3881 | -102.3804
6.3918 | -102.3766
6.3955 | -102.3872
6.3993 | -102.3979
6.403 | -102.4103
6.4067 | -102.4312
6.4104 | -102.4591
6.4142 | -102.5011
6.4179 | -102.5519
6.4216 | -102.6049
6.4253 | -102.6515
6.4291 | -102.6877
6.4328 | -102.7189
6.4365 | -102.7312
6.4403 | -102.7344
6.444 | -102.7159
6.4477 | -102.6833
6.4514 | -102.6235
6.4552 | -102.5819
6.4589 | -102.5354
6.4626 | -102.4858
6.4663 | -102.4437
6.4701 | -102.4059
6.4738 | -102.3701
6.4775 | -102.3477
6.4812 | -102.3364
6.485 | -102.3375
6.4887 | -102.3487
6.4924 | -102.3711
6.4962 | -102.4048
6.4999 | -102.4517
6.5036 | -102.5065
6.5073 | -102.5724
6.5111 | -102.6475
6.5148 | -102.7241
6.5185 | -102.8044
6.5222 | -102.8879
6.526 | -102.969
6.5297 | -103.0384
6.5334 | -103.1034
6.5372 | -103.1566
6.5409 | -103.2037
6.5446 | -103.233
6.5483 | -103.2487
6.5521 | -103.2574
6.5558 | -103.2519
6.5595 | -103.2358
6.5632 | -103.2071
6.567 | -103.1645
6.5707 | -103.1211
6.5744 | -103.0691
6.5781 | -103.012
6.5819 | -102.9564
6.5856 | -102.9136
6.5893 | -102.8781
6.5931 | -102.8485
6.5968 | -102.8277
6.6005 | -102.8123
6.6042 | -102.8027
6.608 | -102.7934
6.6117 | -102.7773
6.6154 | -102.7645
6.6191 | -102.7439
6.6229 | -102.7195
6.6266 | -102.6936
6.6303 | -102.6684
6.6341 | -102.6406
6.6378 | -102.6185
6.6415 | -102.6038
6.6452 | -102.6007
6.649 | -102.6085
6.6527 | -102.6224
6.6564 | -102.6523
6.6601 | -102.6961
6.6639 | -102.7549
6.6676 | -102.8229
6.6713 | -102.8984
6.6751 | -102.9889
6.6788 | -103.0788
6.6825 | -103.166
6.6862 | -103.2399
6.69 | -103.3152
6.6937 | -103.3783
6.6974 | -103.4311
6.7011 | -103.4717
6.7049 | -103.5008
6.7086 | -103.5153
6.7123 | -103.513
6.716 | -103.4931
6.7198 | -103.4594
6.7235 | -103.4101
6.7272 | -103.3493
6.731 | -103.2791
6.7347 | -103.2078
6.7384 | -103.1354
6.7421 | -103.0709
6.7459 | -103.0146
6.7496 | -102.963
6.7533 | -102.9251
6.757 | -102.8914
6.7608 | -102.8674
6.7645 | -102.8525
6.7682 | -102.8479
6.772 | -102.853
6.7757 | -102.8638
6.7794 | -102.8807
6.7831 | -102.9036
6.7869 | -102.9313
6.7906 | -102.956
6.7943 | -102.984
6.798 | -103.0078
6.8018 | -103.0287
6.8055 | -103.0489
6.8092 | -103.0681
6.813 | -103.0855
6.8167 | -103.0943
6.8204 | -103.0972
6.8241 | -103.096
6.8279 | -103.0944
6.8316 | -103.0905
6.8353 | -103.0852
6.839 | -103.0755
6.8428 | -103.0636
6.8465 | -103.0589
6.8502 | -103.0567
6.8539 | -103.0532
6.8577 | -103.0492
6.8614 | -103.0373
6.8651 | -103.0207
6.8689 | -103.0016
6.8726 | -102.9756
6.8763 | -102.9485
6.88 | -102.923
6.8838 | -102.8911
6.8875 | -102.8509
6.8912 | -102.8106
6.8949 | -102.7714
6.8987 | -102.7328
6.9024 | -102.6915
6.9061 | -102.6502
6.9099 | -102.6122
6.9136 | -102.5789
6.9173 | -102.5554
6.921 | -102.541
6.9248 | -102.5366
6.9285 | -102.5424
6.9322 | -102.5657
6.9359 | -102.5998
6.9397 | -102.6501
6.9434 | -102.71
6.9471 | -102.7755
6.9508 | -102.8448
6.9546 | -102.9183
6.9583 | -102.9913
6.962 | -103.0605
6.9658 | -103.1178
6.9695 | -103.1651
6.9732 | -103.2023
6.9769 | -103.2228
6.9807 | -103.2316
6.9844 | -103.2318
6.9881 | -103.1957
6.9918 | -103.1859
6.9956 | -103.161
6.9993 | -103.1245
7.003 | -103.0907
7.0068 | -103.062
7.0105 | -103.0369
7.0142 | -103.0137
7.0179 | -102.9971
7.0217 | -102.9877
7.0254 | -102.9856
7.0291 | -102.9881
7.0328 | -102.9937
7.0366 | -102.9995
7.0403 | -103.015
7.044 | -103.0353
7.0478 | -103.0561
7.0515 | -103.0788
7.0552 | -103.1005
7.0589 | -103.118
7.0627 | -103.1434
7.0664 | -103.174
7.0701 | -103.2057
7.0738 | -103.2342
7.0776 | -103.2633
7.0813 | -103.2965
7.085 | -103.3328
7.0887 | -103.3641
7.0925 | -103.3943
7.0962 | -103.4244
7.0999 | -103.4375
7.1037 | -103.4488
7.1074 | -103.4608
7.1111 | -103.4584
7.1148 | -103.444
7.1186 | -103.4231
7.1223 | -103.3998
7.126 | -103.3699
7.1297 | -103.3325
7.1335 | -103.2961
7.1372 | -103.2554
7.1409 | -103.218
7.1447 | -103.1833
7.1484 | -103.1443
7.1521 | -103.1117
7.1558 | -103.0878
7.1596 | -103.0723
7.1633 | -103.0655
7.167 | -103.0664
7.1707 | -103.0657
7.1745 | -103.0747
7.1782 | -103.0966
7.1819 | -103.1247
7.1857 | -103.153
7.1894 | -103.1833
7.1931 | -103.2109
7.1968 | -103.2425
7.2006 | -103.2748
7.2043 | -103.3052
7.208 | -103.3356
7.2117 | -103.3528
7.2155 | -103.3639
7.2192 | -103.3717
7.2229 | -103.3756
7.2266 | -103.3729
7.2304 | -103.3688
7.2341 | -103.3586
7.2378 | -103.3481
7.2416 | -103.3424
7.2453 | -103.3413
7.249 | -103.3452
7.2527 | -103.355
7.2565 | -103.3742
7.2602 | -103.3933
7.2639 | -103.4169
7.2676 | -103.4391
7.2714 | -103.4648
7.2751 | -103.4915
7.2788 | -103.5166
7.2826 | -103.536
7.2863 | -103.5567
7.29 | -103.5796
7.2937 | -103.6002
7.2975 | -103.6166
7.3012 | -103.6345
7.3049 | -103.6479
7.3086 | -103.6504
7.3124 | -103.6454
7.3161 | -103.635
7.3198 | -103.6206
7.3236 | -103.5926
7.3273 | -103.557
7.331 | -103.5168
7.3347 | -103.47
7.3385 | -103.4189
7.3422 | -103.373
7.3459 | -103.326
7.3496 | -103.2767
7.3534 | -103.2353
7.3571 | -103.1999
7.3608 | -103.166
7.3645 | -103.1428
7.3683 | -103.1316
7.372 | -103.1302
7.3757 | -103.1417
7.3795 | -103.1592
7.3832 | -103.1862
7.3869 | -103.2202
7.3906 | -103.2605
7.3944 | -103.2991
7.3981 | -103.3358
7.4018 | -103.3763
7.4055 | -103.4107
7.4093 | -103.4374
7.413 | -103.4568
7.4167 | -103.4747
7.4205 | -103.4864
7.4242 | -103.4949
7.4279 | -103.4934
7.4316 | -103.4912
7.4354 | -103.494
7.4391 | -103.4982
7.4428 | -103.505
7.4465 | -103.518
7.4503 | -103.5366
7.454 | -103.5649
7.4577 | -103.601
7.4614 | -103.6435
7.4652 | -103.6843
7.4689 | -103.733
7.4726 | -103.7814
7.4764 | -103.8229
7.4801 | -103.8589
7.4838 | -103.8842
7.4875 | -103.9002
7.4913 | -103.9072
7.495 | -103.9062
7.4987 | -103.896
7.5024 | -103.8779
7.5062 | -103.851
7.5099 | -103.8151
7.5136 | -103.7802
7.5174 | -103.7489
7.5211 | -103.7201
7.5248 | -103.7015
7.5285 | -103.6752
7.5323 | -103.6521
7.536 | -103.6334
7.5397 | -103.6215
7.5434 | -103.6101
7.5472 | -103.5966
7.5509 | -103.5811
7.5546 | -103.5627
7.5584 | -103.5393
7.5621 | -103.5111
7.5658 | -103.4735
7.5695 | -103.4277
7.5733 | -103.3748
7.577 | -103.3216
7.5807 | -103.2609
7.5844 | -103.1999
7.5882 | -103.1423
7.5919 | -103.0883
7.5956 | -103.0433
7.5993 | -103.0061
7.6031 | -102.9742
7.6068 | -102.9565
7.6105 | -102.9566
7.6143 | -102.9715
7.618 | -103.0006
7.6217 | -103.0406
7.6254 | -103.0937
7.6292 | -103.1571
7.6329 | -103.2315
7.6366 | -103.3135
7.6403 | -103.401
7.6441 | -103.4869
7.6478 | -103.5709
7.6515 | -103.6493
7.6553 | -103.7254
7.659 | -103.7964
7.6627 | -103.8632
7.6664 | -103.9245
7.6702 | -103.9787
7.6739 | -104.0212
7.6776 | -104.052
7.6813 | -104.081
7.6851 | -104.1098
7.6888 | -104.1278
7.6925 | -104.1318
7.6963 | -104.127
7.7 | -104.1097
7.7037 | -104.0926
7.7074 | -104.0705
7.7112 | -104.0437
7.7149 | -104.018
7.7186 | -103.9943
7.7223 | -103.9687
7.7261 | -103.9345
7.7298 | -103.9014
7.7335 | -103.8719
7.7372 | -103.8409
7.741 | -103.8019
7.7447 | -103.7588
7.7484 | -103.7162
7.7522 | -103.6745
7.7559 | -103.6321
7.7596 | -103.5913
7.7633 | -103.5541
7.7671 | -103.5188
7.7708 | -103.4796
7.7745 | -103.4476
7.7782 | -103.425
7.782 | -103.4116
7.7857 | -103.4048
7.7894 | -103.402
7.7932 | -103.407
7.7969 | -103.4144
7.8006 | -103.4235
7.8043 | -103.4376
7.8081 | -103.4581
7.8118 | -103.4824
7.8155 | -103.5105
7.8192 | -103.5362
7.823 | -103.5618
7.8267 | -103.5934
7.8304 | -103.6279
7.8341 | -103.662
7.8379 | -103.6971
7.8416 | -103.7329
7.8453 | -103.7661
7.8491 | -103.7913
7.8528 | -103.8124
7.8565 | -103.8309
7.8602 | -103.8439
7.864 | -103.8508
7.8677 | -103.8502
7.8714 | -103.8454
7.8751 | -103.8385
7.8789 | -103.8314
7.8826 | -103.8262
7.8863 | -103.8301
7.8901 | -103.8318
7.8938 | -103.8388
7.8975 | -103.8511
7.9012 | -103.8651
7.905 | -103.8876
7.9087 | -103.9153
7.9124 | -103.944
7.9161 | -103.9755
7.9199 | -104.0074
7.9236 | -104.0385
7.9273 | -104.0651
7.9311 | -104.0866
7.9348 | -104.1112
7.9385 | -104.1335
7.9422 | -104.1545
7.946 | -104.1739
7.9497 | -104.1886
7.9534 | -104.1977
7.9571 | -104.2086
7.9609 | -104.2181
7.9646 | -104.2279
7.9683 | -104.2354
7.972 | -104.2427
7.9758 | -104.2511
7.9795 | -104.2545
7.9832 | -104.2612
7.987 | -104.2674
7.9907 | -104.2667
7.9944 | -104.2689
7.9981 | -104.2737
8.0019 | -104.2724
8.0056 | -104.2739
8.0093 | -104.2678
8.013 | -104.2548
8.0168 | -104.2408
8.0205 | -104.2182
8.0242 | -104.19
8.028 | -104.1586
8.0317 | -104.1215
8.0354 | -104.0844
8.0391 | -104.0439
8.0429 | -104.0049
8.0466 | -103.9665
8.0503 | -103.9297
8.054 | -103.895
8.0578 | -103.8628
8.0615 | -103.8367
8.0652 | -103.841
8.0689 | -103.8398
8.0727 | -103.8447
8.0764 | -103.8526
8.0801 | -103.8616
8.0839 | -103.8755
8.0876 | -103.8922
8.0913 | -103.9105
8.095 | -103.9283
8.0988 | -103.9466
8.1025 | -103.9656
8.1062 | -103.9802
8.1099 | -103.9895
8.1137 | -103.9885
8.1174 | -103.9827
8.1211 | -103.9743
8.1249 | -103.9593
8.1286 | -103.9424
8.1323 | -103.9281
8.136 | -103.916
8.1398 | -103.8975
8.1435 | -103.8782
8.1472 | -103.8612
8.1509 | -103.8481
8.1547 | -103.8411
8.1584 | -103.8423
8.1621 | -103.8483
8.1659 | -103.8579
8.1696 | -103.871
8.1733 | -103.8933
8.177 | -103.9142
8.1808 | -103.933
8.1845 | -103.9522
8.1882 | -103.9693
8.1919 | -103.9834
8.1957 | -103.9959
8.1994 | -104.0061
8.2031 | -104.0138
8.2068 | -104.0134
8.2106 | -104.0108
8.2143 | -104.008
8.218 | -104.0029
8.2218 | -103.9972
8.2255 | -103.9849
8.2292 | -103.9723
8.2329 | -103.9641
8.2367 | -103.9539
8.2404 | -103.9441
8.2441 | -103.9389
8.2478 | -103.9323
8.2516 | -103.9257
8.2553 | -103.9156
8.259 | -103.9062
8.2628 | -103.9038
8.2665 | -103.9024
8.2702 | -103.9038
8.2739 | -103.9032
8.2777 | -103.9067
8.2814 | -103.9133
8.2851 | -103.9252
8.2888 | -103.9429
8.2926 | -103.9648
8.2963 | -103.993
8.3 | -104.0258
8.3038 | -104.0644
8.3075 | -104.1032
8.3112 | -104.1425
8.3149 | -104.1852
8.3187 | -104.2255
8.3224 | -104.2599
8.3261 | -104.2955
8.3298 | -104.3295
8.3336 | -104.3552
8.3373 | -104.3726
8.341 | -104.3855
8.3447 | -104.3988
8.3485 | -104.4091
8.3522 | -104.4128
8.3559 | -104.4079
8.3597 | -104.4048
8.3634 | -104.3936
8.3671 | -104.3787
8.3708 | -104.3692
8.3746 | -104.3631
8.3783 | -104.36
8.382 | -104.3503
8.3857 | -104.3399
8.3895 | -104.3323
8.3932 | -104.3277
8.3969 | -104.3253
8.4007 | -104.3124
8.4044 | -104.3023
8.4081 | -104.2886
8.4118 | -104.2731
8.4156 | -104.2572
8.4193 | -104.2381
8.423 | -104.223
8.4267 | -104.2066
8.4305 | -104.1866
8.4342 | -104.1685
8.4379 | -104.1478
8.4416 | -104.1259
8.4454 | -104.1126
8.4491 | -104.0987
8.4528 | -104.0822
8.4566 | -104.0661
8.4603 | -104.0487
8.464 | -104.0302
8.4677 | -104.0042
8.4715 | -103.9773
8.4752 | -103.9512
8.4789 | -103.9251
8.4826 | -103.8968
8.4864 | -103.8722
8.4901 | -103.8486
8.4938 | -103.8221
8.4976 | -103.8017
8.5013 | -103.7917
8.505 | -103.7849
8.5087 | -103.7809
8.5125 | -103.7862
8.5162 | -103.7947
8.5199 | -103.8051
8.5236 | -103.823
8.5274 | -103.8367
8.5311 | -103.8504
8.5348 | -103.8618
8.5386 | -103.8722
8.5423 | -103.881
8.546 | -103.8824
8.5497 | -103.8808
8.5535 | -103.8802
8.5572 | -103.8789
8.5609 | -103.8765
8.5646 | -103.8734
8.5684 | -103.8655
8.5721 | -103.8562
8.5758 | -103.8525
8.5795 | -103.8521
8.5833 | -103.8542
8.587 | -103.8605
8.5907 | -103.8691
8.5945 | -103.8792
8.5982 | -103.8944
8.6019 | -103.9215
8.6056 | -103.9416
8.6094 | -103.9636
8.6131 | -103.99
8.6168 | -104.0167
8.6205 | -104.0467
8.6243 | -104.0744
8.628 | -104.0968
8.6317 | -104.1229
8.6355 | -104.1474
8.6392 | -104.161
8.6429 | -104.1717
8.6466 | -104.1794
8.6504 | -104.1814
8.6541 | -104.185
8.6578 | -104.1885
8.6615 | -104.1879
8.6653 | -104.1828
8.669 | -104.1744
8.6727 | -104.1708
8.6765 | -104.1667
8.6802 | -104.1589
8.6839 | -104.1471
8.6876 | -104.1377
8.6914 | -104.1329
8.6951 | -104.1269
8.6988 | -104.125
8.7025 | -104.1242
8.7063 | -104.1252
8.71 | -104.1352
8.7137 | -104.1497
8.7174 | -104.1677
8.7212 | -104.1893
8.7249 | -104.2097
8.7286 | -104.235
8.7324 | -104.2627
8.7361 | -104.2874
8.7398 | -104.3115
8.7435 | -104.3355
8.7473 | -104.3582
8.751 | -104.379
8.7547 | -104.3948
8.7584 | -104.4035
8.7622 | -104.4064
8.7659 | -104.4055
8.7696 | -104.401
8.7734 | -104.3926
8.7771 | -104.383
8.7808 | -104.3678
8.7845 | -104.3435
8.7883 | -104.3198
8.792 | -104.2977
8.7957 | -104.2738
8.7994 | -104.253
8.8032 | -104.2299
8.8069 | -104.2044
8.8106 | -104.1835
8.8143 | -104.1648
8.8181 | -104.1501
8.8218 | -104.133
8.8255 | -104.1163
8.8293 | -104.0995
8.833 | -104.0865
8.8367 | -104.0745
8.8404 | -104.0638
8.8442 | -104.0607
8.8479 | -104.0608
8.8516 | -104.0638
8.8553 | -104.0735
8.8591 | -104.0806
8.8628 | -104.0894
8.8665 | -104.0998
8.8703 | -104.1122
8.874 | -104.1231
8.8777 | -104.1354
8.8814 | -104.1449
8.8852 | -104.1598
8.8889 | -104.1796
8.8926 | -104.2003
8.8963 | -104.2208
8.9001 | -104.2426
8.9038 | -104.2644
8.9075 | -104.2925
8.9113 | -104.322
8.915 | -104.3487
8.9187 | -104.3765
8.9224 | -104.4104
8.9262 | -104.4474
8.9299 | -104.4827
8.9336 | -104.5195
8.9373 | -104.5539
8.9411 | -104.5836
8.9448 | -104.6118
8.9485 | -104.6339
8.9522 | -104.6526
8.956 | -104.6676
8.9597 | -104.6767
8.9634 | -104.6747
8.9672 | -104.6656
8.9709 | -104.6444
8.9746 | -104.6126
8.9783 | -104.5784
8.9821 | -104.54
8.9858 | -104.4901
8.9895 | -104.4383
8.9932 | -104.3825
8.997 | -104.326
9.0007 | -104.2732
9.0044 | -104.224
9.0082 | -104.1776
9.0119 | -104.1426
9.0156 | -104.1121
9.0193 | -104.0904
9.0231 | -104.0797
9.0268 | -104.078
9.0305 | -104.0852
9.0342 | -104.0984
9.038 | -104.1141
9.0417 | -104.1418
9.0454 | -104.1773
9.0491 | -104.2099
9.0529 | -104.2403
9.0566 | -104.2693
9.0603 | -104.2977
9.0641 | -104.3255
9.0678 | -104.3542
9.0715 | -104.3755
9.0752 | -104.3952
9.079 | -104.4109
9.0827 | -104.4208
9.0864 | -104.4231
9.0901 | -104.4151
9.0939 | -104.4051
9.0976 | -104.3993
9.1013 | -104.389
9.1051 | -104.3768
9.1088 | -104.3618
9.1125 | -104.3421
9.1162 | -104.3281
9.12 | -104.3139
9.1237 | -104.2997
9.1274 | -104.2868
9.1311 | -104.2687
9.1349 | -104.2553
9.1386 | -104.2511
9.1423 | -104.2466
9.1461 | -104.2406
9.1498 | -104.229
9.1535 | -104.2174
9.1572 | -104.2144
9.161 | -104.2159
9.1647 | -104.2164
9.1684 | -104.2207
9.1721 | -104.2283
9.1759 | -104.2319
9.1796 | -104.2357
9.1833 | -104.238
9.187 | -104.2382
9.1908 | -104.2406
9.1945 | -104.2406
9.1982 | -104.234
9.202 | -104.2248
9.2057 | -104.2188
9.2094 | -104.214
9.2131 | -104.2095
9.2169 | -104.2086
9.2206 | -104.2063
9.2243 | -104.205
9.228 | -104.2021
9.2318 | -104.1958
9.2355 | -104.1942
9.2392 | -104.1937
9.243 | -104.1935
9.2467 | -104.1938
9.2504 | -104.1891
9.2541 | -104.1831
9.2579 | -104.1781
9.2616 | -104.171
9.2653 | -104.1634
9.269 | -104.1558
9.2728 | -104.1499
9.2765 | -104.1459
9.2802 | -104.1421
9.284 | -104.1414
9.2877 | -104.146
9.2914 | -104.1518
9.2951 | -104.1486
9.2989 | -104.1453
9.3026 | -104.1478
9.3063 | -104.1511
9.31 | -104.1493
9.3138 | -104.15
9.3175 | -104.153
9.3212 | -104.1601
9.3249 | -104.168
9.3287 | -104.1769
9.3324 | -104.1891
9.3361 | -104.2035
9.3399 | -104.2231
9.3436 | -104.2413
9.3473 | -104.2579
9.351 | -104.2805
9.3548 | -104.3057
9.3585 | -104.3358
9.3622 | -104.372
9.3659 | -104.4002
9.3697 | -104.4255
9.3734 | -104.4516
9.3771 | -104.4773
9.3809 | -104.503
9.3846 | -104.5302
9.3883 | -104.5565
9.392 | -104.5774
9.3958 | -104.5998
9.3995 | -104.6181
9.4032 | -104.6324
9.4069 | -104.6433
9.4107 | -104.6507
9.4144 | -104.6558
9.4181 | -104.6577
9.4218 | -104.6619
9.4256 | -104.6705
9.4293 | -104.6755
9.433 | -104.6838
9.4368 | -104.6964
9.4405 | -104.712
9.4442 | -104.7273
9.4479 | -104.7496
9.4517 | -104.7747
9.4554 | -104.7979
9.4591 | -104.8225
9.4628 | -104.852
9.4666 | -104.8784
9.4703 | -104.904
9.474 | -104.9308
9.4778 | -104.954
9.4815 | -104.9786
9.4852 | -104.9997
9.4889 | -105.0172
9.4927 | -105.0313
9.4964 | -105.0437
9.5001 | -105.0493
9.5038 | -105.0561
9.5076 | -105.0618
9.5113 | -105.0586
9.515 | -105.0463
9.5188 | -105.038
9.5225 | -105.0226
9.5262 | -105.0045
9.5299 | -104.9871
9.5337 | -104.9648
9.5374 | -104.9462
9.5411 | -104.9287
9.5448 | -104.9076
9.5486 | -104.8906
9.5523 | -104.8777
9.556 | -104.8665
9.5597 | -104.8541
9.5635 | -104.8423
9.5672 | -104.8403
9.5709 | -104.8414
9.5747 | -104.8439
9.5784 | -104.8509
9.5821 | -104.8654
9.5858 | -104.881
9.5896 | -104.902
9.5933 | -104.9269
9.597 | -104.9533
9.6007 | -104.9824
9.6045 | -105.0174
9.6082 | -105.051
9.6119 | -105.0883
9.6157 | -105.1201
9.6194 | -105.1471
9.6231 | -105.1798
9.6268 | -105.2059
9.6306 | -105.228
9.6343 | -105.2418
9.638 | -105.2563
9.6417 | -105.2691
9.6455 | -105.2754
9.6492 | -105.2839
9.6529 | -105.2934
9.6567 | -105.2885
9.6604 | -105.2834
9.6641 | -105.275
9.6678 | -105.2602
9.6716 | -105.2378
9.6753 | -105.1955
9.679 | -105.1588
9.6827 | -105.1203
9.6865 | -105.0776
9.6902 | -105.0273
9.6939 | -104.9719
9.6976 | -104.9142
9.7014 | -104.8584
9.7051 | -104.8013
9.7088 | -104.7445
9.7126 | -104.6919
9.7163 | -104.6466
9.72 | -104.606
9.7237 | -104.5714
9.7275 | -104.545
9.7312 | -104.5272
9.7349 | -104.5176
9.7386 | -104.5143
9.7424 | -104.516
9.7461 | -104.5242
9.7498 | -104.5431
9.7536 | -104.566
9.7573 | -104.5977
9.761 | -104.6359
9.7647 | -104.6747
9.7685 | -104.719
9.7722 | -104.7684
9.7759 | -104.8231
9.7796 | -104.8736
9.7834 | -104.92
9.7871 | -104.9619
9.7908 | -105.0001
9.7945 | -105.037
9.7983 | -105.0714
9.802 | -105.0989
9.8057 | -105.1167
9.8095 | -105.1289
9.8132 | -105.1385
9.8169 | -105.1355
9.8206 | -105.1297
9.8244 | -105.1266
9.8281 | -105.116
9.8318 | -105.1024
9.8355 | -105.0876
9.8393 | -105.0731
9.843 | -105.0575
9.8467 | -105.0391
9.8505 | -105.0249
9.8542 | -105.0101
9.8579 | -104.9943
9.8616 | -104.9782
9.8654 | -104.9685
9.8691 | -104.9577
9.8728 | -104.9456
9.8765 | -104.9336
9.8803 | -104.9239
9.884 | -104.9125
9.8877 | -104.8995
9.8915 | -104.8909
9.8952 | -104.8807
9.8989 | -104.8666
9.9026 | -104.8563
9.9064 | -104.8461
9.9101 | -104.836
9.9138 | -104.8312
9.9175 | -104.8263
9.9213 | -104.8227
9.925 | -104.8265
9.9287 | -104.8375
9.9324 | -104.8526
9.9362 | -104.8691
9.9399 | -104.8858
9.9436 | -104.9038
9.9474 | -104.9276
9.9511 | -104.9575
9.9548 | -104.9848
9.9585 | -105.0078
9.9623 | -105.0366
9.966 | -105.0631
9.9697 | -105.0855
9.9734 | -105.1077
9.9772 | -105.1271
9.9809 | -105.1416
9.9846 | -105.1488
9.9884 | -105.1533
9.9921 | -105.1531
9.9958 | -105.144
9.9995 | -105.1345
10.0033 | -105.1251
10.007 | -105.1091
10.0107 | -105.0919
10.0144 | -105.0776
10.0182 | -105.0637
10.0219 | -105.0437
10.0256 | -105.0269
10.0294 | -105.0154
10.0331 | -105.0067
10.0368 | -105.0002
10.0405 | -104.9995
10.0443 | -105.0059
10.048 | -105.0097
10.0517 | -105.0117
10.0554 | -105.0134
10.0592 | -105.0137
10.0629 | -105.0156
10.0666 | -105.0203
10.0703 | -105.0197
10.0741 | -105.0128
10.0778 | -105.0093
10.0815 | -105.0022
10.0853 | -104.9913
10.089 | -104.9793
10.0927 | -104.9663
10.0964 | -104.9535
10.1002 | -104.9419
10.1039 | -104.9314
10.1076 | -104.9169
10.1113 | -104.8973
10.1151 | -104.8761
10.1188 | -104.8492
10.1225 | -104.8259
10.1263 | -104.804
10.13 | -104.7821
10.1337 | -104.759
10.1374 | -104.7374
10.1412 | -104.7156
10.1449 | -104.6924
10.1486 | -104.673
10.1523 | -104.6599
10.1561 | -104.6465
10.1598 | -104.6378
10.1635 | -104.6346
10.1672 | -104.6332
10.171 | -104.6325
10.1747 | -104.6359
10.1784 | -104.643
10.1822 | -104.6509
10.1859 | -104.6599
10.1896 | -104.6744
10.1933 | -104.693
10.1971 | -104.7169
10.2008 | -104.7467
10.2045 | -104.7753
10.2082 | -104.8045
10.212 | -104.8392
10.2157 | -104.8841
10.2194 | -104.9208
10.2232 | -104.9611
10.2269 | -105.0006
10.2306 | -105.0413
10.2343 | -105.0815
10.2381 | -105.121
10.2418 | -105.1526
10.2455 | -105.1822
10.2492 | -105.2126
10.253 | -105.2408
10.2567 | -105.267
10.2604 | -105.2926
10.2642 | -105.3162
10.2679 | -105.3398
10.2716 | -105.3583
10.2753 | -105.3758
10.2791 | -105.3907
10.2828 | -105.4029
10.2865 | -105.4143
10.2902 | -105.4261
10.294 | -105.4365
10.2977 | -105.4432
10.3014 | -105.4438
10.3051 | -105.4472
10.3089 | -105.4501
10.3126 | -105.4511
10.3163 | -105.4567
10.3201 | -105.4546
10.3238 | -105.4408
10.3275 | -105.4265
10.3312 | -105.4139
10.335 | -105.4058
10.3387 | -105.395
10.3424 | -105.3853
10.3461 | -105.3753
10.3499 | -105.3589
10.3536 | -105.3467
10.3573 | -105.3357
10.3611 | -105.3202
10.3648 | -105.3071
10.3685 | -105.2999
10.3722 | -105.2929
10.376 | -105.2834
10.3797 | -105.2713
10.3834 | -105.2649
10.3871 | -105.2636
10.3909 | -105.2616
10.3946 | -105.2608
10.3983 | -105.2577
10.402 | -105.2482
10.4058 | -105.2353
10.4095 | -105.2243
10.4132 | -105.2163
10.417 | -105.2094
10.4207 | -105.201
10.4244 | -105.1882
10.4281 | -105.1713
10.4319 | -105.1576
10.4356 | -105.1418
10.4393 | -105.1212
10.443 | -105.0967
10.4468 | -105.0732
10.4505 | -105.0507
10.4542 | -105.0305
10.458 | -105.005
10.4617 | -104.9794
10.4654 | -104.9578
10.4691 | -104.941
10.4729 | -104.9301
10.4766 | -104.9258
10.4803 | -104.92
10.484 | -104.9142
10.4878 | -104.9074
10.4915 | -104.9091
10.4952 | -104.915
10.499 | -104.9207
10.5027 | -104.9295
10.5064 | -104.9368
10.5101 | -104.9443
10.5139 | -104.9516
10.5176 | -104.9566
10.5213 | -104.9522
10.525 | -104.9485
10.5288 | -104.9456
10.5325 | -104.9366
10.5362 | -104.9203
10.5399 | -104.8971
10.5437 | -104.8744
10.5474 | -104.8492
10.5511 | -104.8249
10.5549 | -104.804
10.5586 | -104.7809
10.5623 | -104.7576
10.566 | -104.7316
10.5698 | -104.707
10.5735 | -104.6902
10.5772 | -104.678
10.5809 | -104.6715
10.5847 | -104.6694
10.5884 | -104.6737
10.5921 | -104.6818
10.5959 | -104.6975
10.5996 | -104.7165
10.6033 | -104.7372
10.607 | -104.7634
10.6108 | -104.7929
10.6145 | -104.8239
10.6182 | -104.8536
10.6219 | -104.8837
10.6257 | -104.9117
10.6294 | -104.94
10.6331 | -104.9652
10.6369 | -104.9849
10.6406 | -105.0064
10.6443 | -105.0229
10.648 | -105.0362
10.6518 | -105.0477
10.6555 | -105.0556
10.6592 | -105.0618
10.6629 | -105.0632
10.6667 | -105.0625
10.6704 | -105.0627
10.6741 | -105.0665
10.6778 | -105.0678
10.6816 | -105.0661
10.6853 | -105.0664
10.689 | -105.067
10.6928 | -105.0696
10.6965 | -105.0705
10.7002 | -105.0745
10.7039 | -105.0781
10.7077 | -105.0839
10.7114 | -105.0952
10.7151 | -105.1059
10.7188 | -105.1188
10.7226 | -105.1343
10.7263 | -105.1482
10.73 | -105.1629
10.7338 | -105.1805
10.7375 | -105.1988
10.7412 | -105.2164
10.7449 | -105.2344
10.7487 | -105.2494
10.7524 | -105.262
10.7561 | -105.2716
10.7598 | -105.2827
10.7636 | -105.2897
10.7673 | -105.2992
10.771 | -105.3083
10.7747 | -105.3154
10.7785 | -105.3162
10.7822 | -105.318
10.7859 | -105.3177
10.7897 | -105.317
10.7934 | -105.3158
10.7971 | -105.3118
10.8008 | -105.3057
10.8046 | -105.2996
10.8083 | -105.2965
10.812 | -105.294
10.8157 | -105.2902
10.8195 | -105.2856
10.8232 | -105.2808
10.8269 | -105.2757
10.8307 | -105.2713
10.8344 | -105.266
10.8381 | -105.262
10.8418 | -105.2587
10.8456 | -105.2569
10.8493 | -105.2555
10.853 | -105.2535
10.8567 | -105.2555
10.8605 | -105.2611
10.8642 | -105.2656
10.8679 | -105.268
10.8717 | -105.2695
10.8754 | -105.2753
10.8791 | -105.2821
10.8828 | -105.2908
10.8866 | -105.3
10.8903 | -105.3078
10.894 | -105.3154
10.8977 | -105.3233
10.9015 | -105.3296
10.9052 | -105.3331
10.9089 | -105.3315
10.9126 | -105.3331
10.9164 | -105.3375
10.9201 | -105.3356
10.9238 | -105.3317
10.9276 | -105.3228
10.9313 | -105.3132
10.935 | -105.3062
10.9387 | -105.2992
10.9425 | -105.2916
10.9462 | -105.2829
10.9499 | -105.2721
10.9536 | -105.2615
10.9574 | -105.2489
10.9611 | -105.2389
10.9648 | -105.2294
10.9686 | -105.2223
10.9723 | -105.2195
10.976 | -105.2182
10.9797 | -105.223
10.9835 | -105.2281
10.9872 | -105.2336
10.9909 | -105.2425
10.9946 | -105.257
10.9984 | -105.2734
11.0021 | -105.2902
11.0058 | -105.3082
11.0096 | -105.3303
11.0133 | -105.356
11.017 | -105.3816
11.0207 | -105.4081
11.0245 | -105.4354
11.0282 | -105.4628
11.0319 | -105.4899
11.0356 | -105.5121
11.0394 | -105.5317
11.0431 | -105.5502
11.0468 | -105.5675
11.0505 | -105.5808
11.0543 | -105.5909
11.058 | -105.5994
11.0617 | -105.6077
11.0655 | -105.6146
11.0692 | -105.6215
11.0729 | -105.6172
11.0766 | -105.6099
11.0804 | -105.6064
11.0841 | -105.6047
11.0878 | -105.6008
11.0915 | -105.5921
11.0953 | -105.5774
11.099 | -105.5659
11.1027 | -105.5477
11.1065 | -105.5272
11.1102 | -105.5057
11.1139 | -105.4775
11.1176 | -105.4461
11.1214 | -105.4142
11.1251 | -105.3835
11.1288 | -105.3491
11.1325 | -105.3114
11.1363 | -105.2756
11.14 | -105.2412
11.1437 | -105.2091
11.1474 | -105.18
11.1512 | -105.1537
11.1549 | -105.1308
11.1586 | -105.1085
11.1624 | -105.0901
11.1661 | -105.0777
11.1698 | -105.0663
11.1735 | -105.0575
11.1773 | -105.0509
11.181 | -105.0452
11.1847 | -105.0463
11.1884 | -105.0495
11.1922 | -105.0566
11.1959 | -105.0644
11.1996 | -105.0733
11.2034 | -105.0869
11.2071 | -105.1013
11.2108 | -105.1156
11.2145 | -105.1319
11.2183 | -105.1519
11.222 | -105.1717
11.2257 | -105.1894
11.2294 | -105.2082
11.2332 | -105.2286
11.2369 | -105.2478
11.2406 | -105.2616
11.2444 | -105.274
11.2481 | -105.2846
11.2518 | -105.2935
11.2555 | -105.3026
11.2593 | -105.3085
11.263 | -105.3161
11.2667 | -105.3187
11.2704 | -105.3177
11.2742 | -105.3171
11.2779 | -105.3144
11.2816 | -105.3064
11.2853 | -105.2984
11.2891 | -105.2822
11.2928 | -105.2734
11.2965 | -105.2662
11.3003 | -105.2582
11.304 | -105.2461
11.3077 | -105.234
11.3114 | -105.2207
11.3152 | -105.2086
11.3189 | -105.1967
11.3226 | -105.1888
11.3263 | -105.1822
11.3301 | -105.1781
11.3338 | -105.1761
11.3375 | -105.177
11.3413 | -105.1829
11.345 | -105.1925
11.3487 | -105.2021
11.3524 | -105.2152
11.3562 | -105.2302
11.3599 | -105.2514
11.3636 | -105.2744
11.3673 | -105.2973
11.3711 | -105.325
11.3748 | -105.3543
11.3785 | -105.3811
11.3823 | -105.4076
11.386 | -105.4367
11.3897 | -105.4675
11.3934 | -105.4953
11.3972 | -105.5209
11.4009 | -105.5453
11.4046 | -105.5728
11.4083 | -105.5946
11.4121 | -105.616
11.4158 | -105.6365
11.4195 | -105.652
11.4232 | -105.6605
11.427 | -105.6686
11.4307 | -105.6777
11.4344 | -105.6868
11.4382 | -105.6861
11.4419 | -105.682
11.4456 | -105.6788
11.4493 | -105.6715
11.4531 | -105.6629
11.4568 | -105.655
11.4605 | -105.6448
11.4642 | -105.6334
11.468 | -105.6223
11.4717 | -105.6149
11.4754 | -105.609
11.4792 | -105.6043
11.4829 | -105.5985
11.4866 | -105.5915
11.4903 | -105.5864
11.4941 | -105.5855
11.4978 | -105.5849
11.5015 | -105.5863
11.5052 | -105.5893
11.509 | -105.592
11.5127 | -105.5981
11.5164 | -105.6042
11.5201 | -105.6072
11.5239 | -105.6109
11.5276 | -105.6105
11.5313 | -105.6057
11.5351 | -105.5999
11.5388 | -105.5934
11.5425 | -105.5867
11.5462 | -105.5763
11.55 | -105.5671
11.5537 | -105.5553
11.5574 | -105.542
11.5611 | -105.5287
11.5649 | -105.5149
11.5686 | -105.4997
11.5723 | -105.4811
11.5761 | -105.4635
11.5798 | -105.4466
11.5835 | -105.4271
11.5872 | -105.4129
11.591 | -105.4025
11.5947 | -105.3903
11.5984 | -105.3771
11.6021 | -105.3677
11.6059 | -105.3649
11.6096 | -105.3602
11.6133 | -105.3579
11.6171 | -105.3546
11.6208 | -105.3518
11.6245 | -105.3528
11.6282 | -105.3549
11.632 | -105.3587
11.6357 | -105.3655
11.6394 | -105.3705
11.6431 | -105.376
11.6469 | -105.3834
11.6506 | -105.3892
11.6543 | -105.3971
11.658 | -105.4052
11.6618 | -105.4112
11.6655 | -105.4139
11.6692 | -105.4204
11.673 | -105.427
11.6767 | -105.4277
11.6804 | -105.428
11.6841 | -105.4294
11.6879 | -105.434
11.6916 | -105.4329
11.6953 | -105.4291
11.699 | -105.4286
11.7028 | -105.4286
11.7065 | -105.4238
11.7102 | -105.4175
11.714 | -105.4126
11.7177 | -105.4048
11.7214 | -105.3959
11.7251 | -105.3876
11.7289 | -105.3777
11.7326 | -105.3691
11.7363 | -105.3609
11.74 | -105.3561
11.7438 | -105.3497
11.7475 | -105.3434
11.7512 | -105.341
11.7549 | -105.3351
11.7587 | -105.3285
11.7624 | -105.3273
11.7661 | -105.3328
11.7699 | -105.3342
11.7736 | -105.3382
11.7773 | -105.3462
11.781 | -105.3518
11.7848 | -105.3624
11.7885 | -105.3754
11.7922 | -105.3826
11.7959 | -105.3927
11.7997 | -105.4061
11.8034 | -105.4188
11.8071 | -105.4331
11.8109 | -105.4467
11.8146 | -105.4596
11.8183 | -105.4739
11.822 | -105.4852
11.8258 | -105.4852
11.8295 | -105.491
11.8332 | -105.4977
11.8369 | -105.5056
11.8407 | -105.5109
11.8444 | -105.5165
11.8481 | -105.5229
11.8519 | -105.5266
11.8556 | -105.5285
11.8593 | -105.5259
11.863 | -105.5244
11.8668 | -105.524
11.8705 | -105.519
11.8742 | -105.511
11.8779 | -105.5059
11.8817 | -105.5007
11.8854 | -105.4939
11.8891 | -105.4846
11.8928 | -105.4773
11.8966 | -105.4697
11.9003 | -105.4629
11.904 | -105.4546
11.9078 | -105.4456
11.9115 | -105.4368
11.9152 | -105.4313
11.9189 | -105.4256
11.9227 | -105.421
11.9264 | -105.419
11.9301 | -105.4225
11.9338 | -105.4242
11.9376 | -105.4243
11.9413 | -105.4318
11.945 | -105.4425
11.9488 | -105.4538
11.9525 | -105.4656
11.9562 | -105.4802
11.9599 | -105.4955
11.9637 | -105.5167
11.9674 | -105.5395
11.9711 | -105.5598
11.9748 | -105.5817
11.9786 | -105.6048
11.9823 | -105.627
11.986 | -105.6488
11.9898 | -105.6699
11.9935 | -105.6948
11.9972 | -105.721
12.0009 | -105.7447
12.0047 | -105.7654
12.0084 | -105.7857
12.0121 | -105.8025
12.0158 | -105.817
12.0196 | -105.8332
12.0233 | -105.8499
12.027 | -105.8616
12.0307 | -105.8727
12.0345 | -105.8822
12.0382 | -105.8894
12.0419 | -105.8916
12.0457 | -105.8909
12.0494 | -105.8858
12.0531 | -105.8783
12.0568 | -105.8705
12.0606 | -105.8597
12.0643 | -105.8461
12.068 | -105.8249
12.0717 | -105.8015
12.0755 | -105.7824
12.0792 | -105.7636
12.0829 | -105.7455
12.0867 | -105.7251
12.0904 | -105.7075
12.0941 | -105.69
12.0978 | -105.6703
12.1016 | -105.653
12.1053 | -105.6342
12.109 | -105.6164
12.1127 | -105.6004
12.1165 | -105.5859
12.1202 | -105.5735
12.1239 | -105.5618
12.1276 | -105.5514
12.1314 | -105.542
12.1351 | -105.5356
12.1388 | -105.5292
12.1426 | -105.524
12.1463 | -105.52
12.15 | -105.5162
12.1537 | -105.5167
12.1575 | -105.5204
12.1612 | -105.5258
12.1649 | -105.5314
12.1686 | -105.5341
12.1724 | -105.5381
12.1761 | -105.5465
12.1798 | -105.5554
12.1836 | -105.5613
12.1873 | -105.57
12.191 | -105.5812
12.1947 | -105.5885
12.1985 | -105.5944
12.2022 | -105.6032
12.2059 | -105.6072
12.2096 | -105.613
12.2134 | -105.615
12.2171 | -105.619
12.2208 | -105.619
12.2246 | -105.6164
12.2283 | -105.6136
12.232 | -105.6076
12.2357 | -105.6011
12.2395 | -105.5964
12.2432 | -105.593
12.2469 | -105.5854
12.2506 | -105.573
12.2544 | -105.5612
12.2581 | -105.5463
12.2618 | -105.5335
12.2655 | -105.52
12.2693 | -105.5108
12.273 | -105.5056
12.2767 | -105.4962
12.2805 | -105.4877
12.2842 | -105.4826
12.2879 | -105.4779
12.2916 | -105.4725
12.2954 | -105.4664
12.2991 | -105.4658
12.3028 | -105.4695
12.3065 | -105.4723
12.3103 | -105.4767
12.314 | -105.4825
12.3177 | -105.4937
12.3215 | -105.5099
12.3252 | -105.5274
12.3289 | -105.5401
12.3326 | -105.5565
12.3364 | -105.5783
12.3401 | -105.602
12.3438 | -105.628
12.3475 | -105.655
12.3513 | -105.6784
12.355 | -105.6976
12.3587 | -105.7209
12.3625 | -105.7471
12.3662 | -105.7633
12.3699 | -105.7795
12.3736 | -105.7961
12.3774 | -105.8149
12.3811 | -105.8271
12.3848 | -105.8368
12.3885 | -105.8406
12.3923 | -105.8431
12.396 | -105.8411
12.3997 | -105.8386
12.4034 | -105.8363
12.4072 | -105.8268
12.4109 | -105.8135
12.4146 | -105.8002
12.4184 | -105.7838
12.4221 | -105.766
12.4258 | -105.7466
12.4295 | -105.7226
12.4333 | -105.6986
12.437 | -105.6768
12.4407 | -105.6536
12.4444 | -105.628
12.4482 | -105.6003
12.4519 | -105.5727
12.4556 | -105.5497
12.4594 | -105.5261
12.4631 | -105.5026
12.4668 | -105.4836
12.4705 | -105.4679
12.4743 | -105.4553
12.478 | -105.4436
12.4817 | -105.4359
12.4854 | -105.4322
12.4892 | -105.4298
12.4929 | -105.4333
12.4966 | -105.44
12.5003 | -105.4525
12.5041 | -105.4673
12.5078 | -105.4864
12.5115 | -105.5063
12.5153 | -105.5255
12.519 | -105.5499
12.5227 | -105.5782
12.5264 | -105.6076
12.5302 | -105.6397
12.5339 | -105.6746
12.5376 | -105.7058
12.5413 | -105.7341
12.5451 | -105.7647
12.5488 | -105.7951
12.5525 | -105.8251
12.5563 | -105.8546
12.56 | -105.8827
12.5637 | -105.91
12.5674 | -105.9376
12.5712 | -105.9681
12.5749 | -105.994
12.5786 | -106.017
12.5823 | -106.0416
12.5861 | -106.0613
12.5898 | -106.0821
12.5935 | -106.0999
12.5973 | -106.1109
12.601 | -106.1202
12.6047 | -106.1256
12.6084 | -106.1294
12.6122 | -106.1319
12.6159 | -106.13
12.6196 | -106.1256
12.6233 | -106.1175
12.6271 | -106.1119
12.6308 | -106.1047
12.6345 | -106.0914
12.6382 | -106.0734
12.642 | -106.0555
12.6457 | -106.0385
12.6494 | -106.0208
12.6532 | -105.9996
12.6569 | -105.9805
12.6606 | -105.9615
12.6643 | -105.9375
12.6681 | -105.9151
12.6718 | -105.8952
12.6755 | -105.8704
12.6792 | -105.846
12.683 | -105.8213
12.6867 | -105.8032
12.6904 | -105.7902
12.6942 | -105.7758
12.6979 | -105.7602
12.7016 | -105.744
12.7053 | -105.7324
12.7091 | -105.7255
12.7128 | -105.7189
12.7165 | -105.7195
12.7202 | -105.7246
12.724 | -105.7321
12.7277 | -105.7432
12.7314 | -105.7589
12.7352 | -105.776
12.7389 | -105.7942
12.7426 | -105.8125
12.7463 | -105.8323
12.7501 | -105.851
12.7538 | -105.867
12.7575 | -105.8898
12.7612 | -105.9121
12.765 | -105.9279
12.7687 | -105.945
12.7724 | -105.961
12.7761 | -105.9815
12.7799 | -106.0019
12.7836 | -106.018
12.7873 | -106.0335
12.7911 | -106.0441
12.7948 | -106.0524
12.7985 | -106.0563
12.8022 | -106.056
12.806 | -106.0546
12.8097 | -106.0493
12.8134 | -106.0478
12.8171 | -106.0482
12.8209 | -106.0453
12.8246 | -106.0377
12.8283 | -106.0269
12.8321 | -106.0132
12.8358 | -106.0009
12.8395 | -105.9826
12.8432 | -105.9618
12.847 | -105.9468
12.8507 | -105.9289
12.8544 | -105.9088
12.8581 | -105.8911
12.8619 | -105.8697
12.8656 | -105.8463
12.8693 | -105.8225
12.873 | -105.7973
12.8768 | -105.7736
12.8805 | -105.752
12.8842 | -105.7302
12.888 | -105.703
12.8917 | -105.6759
12.8954 | -105.6515
12.8991 | -105.6251
12.9029 | -105.5992
12.9066 | -105.5756
12.9103 | -105.555
12.914 | -105.5353
12.9178 | -105.5154
12.9215 | -105.4981
12.9252 | -105.4825
12.929 | -105.4693
12.9327 | -105.4602
12.9364 | -105.4534
12.9401 | -105.4473
12.9439 | -105.4405
12.9476 | -105.4373
12.9513 | -105.4384
12.955 | -105.44
12.9588 | -105.4423
12.9625 | -105.4468
12.9662 | -105.456
12.97 | -105.4645
12.9737 | -105.4758
12.9774 | -105.4918
12.9811 | -105.5101
12.9849 | -105.5295
12.9886 | -105.5483
12.9923 | -105.5718
12.996 | -105.5951
12.9998 | -105.6162
13.0035 | -105.6374
13.0072 | -105.6591
13.0109 | -105.6794
13.0147 | -105.6976
13.0184 | -105.7112
13.0221 | -105.7256
13.0259 | -105.7417
13.0296 | -105.7546
13.0333 | -105.7593
13.037 | -105.7651
13.0408 | -105.7693
13.0445 | -105.7667
13.0482 | -105.7622
13.0519 | -105.7565
13.0557 | -105.7458
13.0594 | -105.7355
13.0631 | -105.7224
13.0669 | -105.7113
13.0706 | -105.7
13.0743 | -105.6862
13.078 | -105.6727
13.0818 | -105.6593
13.0855 | -105.6475
13.0892 | -105.636
13.0929 | -105.6244
13.0967 | -105.6109
13.1004 | -105.5988
13.1041 | -105.5918
13.1078 | -105.5871
13.1116 | -105.5808
13.1153 | -105.5772
13.119 | -105.5754
13.1228 | -105.5775
13.1265 | -105.5822
13.1302 | -105.5862
13.1339 | -105.5903
13.1377 | -105.5962
13.1414 | -105.6064
13.1451 | -105.6144
13.1488 | -105.6265
13.1526 | -105.6395
13.1563 | -105.6526
13.16 | -105.6695
13.1638 | -105.6864
13.1675 | -105.6997
13.1712 | -105.7126
13.1749 | -105.7265
13.1787 | -105.7405
13.1824 | -105.7465
13.1861 | -105.7497
13.1898 | -105.7529
13.1936 | -105.7584
13.1973 | -105.7653
13.201 | -105.7731
13.2048 | -105.7795
13.2085 | -105.7867
13.2122 | -105.7917
13.2159 | -105.7929
13.2197 | -105.7931
13.2234 | -105.7925
13.2271 | -105.789
13.2308 | -105.7879
13.2346 | -105.7902
13.2383 | -105.7938
13.242 | -105.7948
13.2457 | -105.7965
13.2495 | -105.7993
13.2532 | -105.8004
13.2569 | -105.8007
13.2607 | -105.8043
13.2644 | -105.8124
13.2681 | -105.8186
13.2718 | -105.8234
13.2756 | -105.8269
13.2793 | -105.8312
13.283 | -105.8358
13.2867 | -105.8399
13.2905 | -105.8449
13.2942 | -105.848
13.2979 | -105.8524
13.3017 | -105.8589
13.3054 | -105.8611
13.3091 | -105.8637
13.3128 | -105.8644
13.3166 | -105.8645
13.3203 | -105.8641
13.324 | -105.8634
13.3277 | -105.8661
13.3315 | -105.8626
13.3352 | -105.8639
13.3389 | -105.8614
13.3427 | -105.8556
13.3464 | -105.8499
13.3501 | -105.8477
13.3538 | -105.844
13.3576 | -105.8418
13.3613 | -105.838
13.365 | -105.8331
13.3687 | -105.8299
13.3725 | -105.8287
13.3762 | -105.8265
13.3799 | -105.8237
13.3836 | -105.819
13.3874 | -105.8114
13.3911 | -105.8034
13.3948 | -105.7976
13.3986 | -105.7902
13.4023 | -105.7823
13.406 | -105.7793
13.4097 | -105.7769
13.4135 | -105.7722
13.4172 | -105.7703
13.4209 | -105.7666
13.4246 | -105.7639
13.4284 | -105.7579
13.4321 | -105.7561
13.4358 | -105.7533
13.4396 | -105.748
13.4433 | -105.7403
13.447 | -105.7379
13.4507 | -105.735
13.4545 | -105.7297
13.4582 | -105.7242
13.4619 | -105.7211
13.4656 | -105.7198
13.4694 | -105.7213
13.4731 | -105.723
13.4768 | -105.7259
13.4805 | -105.7268
13.4843 | -105.7274
13.488 | -105.7297
13.4917 | -105.7371
13.4955 | -105.7453
13.4992 | -105.7543
13.5029 | -105.7629
13.5066 | -105.7733
13.5104 | -105.7838
13.5141 | -105.7994
13.5178 | -105.8161
13.5215 | -105.8271
13.5253 | -105.8379
13.529 | -105.8502
13.5327 | -105.8638
13.5365 | -105.8792
13.5402 | -105.8926
13.5439 | -105.9028
13.5476 | -105.912
13.5514 | -105.9187
13.5551 | -105.925
13.5588 | -105.9299
13.5625 | -105.9318
13.5663 | -105.9325
13.57 | -105.9365
13.5737 | -105.9372
13.5775 | -105.9366
13.5812 | -105.9363
13.5849 | -105.9329
13.5886 | -105.9314
13.5924 | -105.9308
13.5961 | -105.9275
13.5998 | -105.9207
13.6035 | -105.9133
13.6073 | -105.9036
13.611 | -105.8915
13.6147 | -105.8797
13.6184 | -105.8648
13.6222 | -105.8508
13.6259 | -105.8338
13.6296 | -105.8129
13.6334 | -105.7913
13.6371 | -105.7697
13.6408 | -105.7513
13.6445 | -105.7317
13.6483 | -105.7091
13.652 | -105.6865
13.6557 | -105.6669
13.6594 | -105.6467
13.6632 | -105.6252
13.6669 | -105.6028
13.6706 | -105.5833
13.6744 | -105.5614
13.6781 | -105.5423
13.6818 | -105.5242
13.6855 | -105.5062
13.6893 | -105.488
13.693 | -105.4717
13.6967 | -105.4602
13.7004 | -105.4511
13.7042 | -105.4404
13.7079 | -105.4302
13.7116 | -105.4216
13.7154 | -105.4153
13.7191 | -105.4131
13.7228 | -105.4143
13.7265 | -105.4149
13.7303 | -105.4195
13.734 | -105.4247
13.7377 | -105.4334
13.7414 | -105.4455
13.7452 | -105.4615
13.7489 | -105.4816
13.7526 | -105.5018
13.7563 | -105.5221
13.7601 | -105.5444
13.7638 | -105.5686
13.7675 | -105.5902
13.7713 | -105.6119
13.775 | -105.6368
13.7787 | -105.6624
13.7824 | -105.6885
13.7862 | -105.7119
13.7899 | -105.7365
13.7936 | -105.7602
13.7973 | -105.7801
13.8011 | -105.8011
13.8048 | -105.8238
13.8085 | -105.8405
13.8123 | -105.8552
13.816 | -105.8688
13.8197 | -105.8792
13.8234 | -105.8889
13.8272 | -105.8994
13.8309 | -105.9089
13.8346 | -105.9152
13.8383 | -105.9178
13.8421 | -105.9185
13.8458 | -105.9206
13.8495 | -105.9224
13.8532 | -105.9233
13.857 | -105.9243
13.8607 | -105.9275
13.8644 | -105.9299
13.8682 | -105.9324
13.8719 | -105.9378
13.8756 | -105.9425
13.8793 | -105.9448
13.8831 | -105.9468
13.8868 | -105.9504
13.8905 | -105.9541
13.8942 | -105.9569
13.898 | -105.9563
13.9017 | -105.9575
13.9054 | -105.96
13.9092 | -105.9604
13.9129 | -105.9603
13.9166 | -105.9576
13.9203 | -105.9566
13.9241 | -105.9519
13.9278 | -105.9444
13.9315 | -105.9368
13.9352 | -105.9293
13.939 | -105.9194
13.9427 | -105.9101
13.9464 | -105.9013
13.9502 | -105.8923
13.9539 | -105.8847
13.9576 | -105.8748
13.9613 | -105.8489
13.9651 | -105.8118
13.9688 | -105.7831
13.9725 | -105.7544
13.9762 | -105.7236
13.98 | -105.6979
13.9837 | -105.6716
13.9874 | -105.6423
13.9911 | -105.6125
13.9949 | -105.5851
13.9986 | -105.5587
14.0023 | -105.5293
14.0061 | -105.4977
14.0098 | -105.4676
14.0135 | -105.4357
14.0172 | -105.405
14.021 | -105.3755
14.0247 | -105.3465
14.0284 | -105.3187
14.0321 | -105.2887
14.0359 | -105.2645
14.0396 | -105.2434
14.0433 | -105.2242
14.0471 | -105.2116
14.0508 | -105.2047
14.0545 | -105.2014
14.0582 | -105.2045
14.062 | -105.2132
14.0657 | -105.2272
14.0694 | -105.2447
14.0731 | -105.2664
14.0769 | -105.2942
14.0806 | -105.3292
14.0843 | -105.3677
14.0881 | -105.4059
14.0918 | -105.4488
14.0955 | -105.495
14.0992 | -105.5433
14.103 | -105.5914
14.1067 | -105.6391
14.1104 | -105.6873
14.1141 | -105.734
14.1179 | -105.7799
14.1216 | -105.8226
14.1253 | -105.8662
14.129 | -105.9049
14.1328 | -105.9412
14.1365 | -105.9739
14.1402 | -106.0037
14.144 | -106.0306
14.1477 | -106.053
14.1514 | -106.072
14.1551 | -106.0903
14.1589 | -106.1067
14.1626 | -106.1311
14.1663 | -106.1594
14.17 | -106.1671
14.1738 | -106.1696
14.1775 | -106.1716
14.1812 | -106.168
14.185 | -106.1651
14.1887 | -106.1591
14.1924 | -106.1496
14.1961 | -106.141
14.1999 | -106.1323
14.2036 | -106.1237
14.2073 | -106.1172
14.211 | -106.1123
14.2148 | -106.1047
14.2185 | -106.0967
14.2222 | -106.0861
14.2259 | -106.0746
14.2297 | -106.0644
14.2334 | -106.0545
14.2371 | -106.048
14.2409 | -106.0376
14.2446 | -106.0294
14.2483 | -106.0224
14.252 | -106.0113
14.2558 | -106.004
14.2595 | -105.9958
14.2632 | -105.9876
14.2669 | -105.981
14.2707 | -105.9783
14.2744 | -105.9769
14.2781 | -105.9776
14.2819 | -105.9777
14.2856 | -105.9759
14.2893 | -105.9768
14.293 | -105.9833
14.2968 | -105.988
14.3005 | -105.9912
14.3042 | -105.9958
14.3079 | -105.9999
14.3117 | -106.0041
14.3154 | -106.0047
14.3191 | -106.0053
14.3229 | -106.0108
14.3266 | -106.0176
14.3303 | -106.0235
14.334 | -106.0323
14.3378 | -106.0393
14.3415 | -106.0439
14.3452 | -106.049
14.3489 | -106.0525
14.3527 | -106.0549
14.3564 | -106.0587
14.3601 | -106.0619
14.3638 | -106.0668
14.3676 | -106.0718
14.3713 | -106.074
14.375 | -106.079
14.3788 | -106.0835
14.3825 | -106.0832
14.3862 | -106.0867
14.3899 | -106.0895
14.3937 | -106.094
14.3974 | -106.1011
14.4011 | -106.1042
14.4048 | -106.1074
14.4086 | -106.1128
14.4123 | -106.1148
14.416 | -106.1156
14.4198 | -106.1194
14.4235 | -106.1235
14.4272 | -106.1316
14.4309 | -106.1389
14.4347 | -106.1421
14.4384 | -106.1446
14.4421 | -106.1455
14.4458 | -106.1456
14.4496 | -106.1452
14.4533 | -106.1474
14.457 | -106.1496
14.4608 | -106.1487
14.4645 | -106.1493
14.4682 | -106.1493
14.4719 | -106.1492
14.4757 | -106.1492
14.4794 | -106.1521
14.4831 | -106.1561
14.4868 | -106.163
14.4906 | -106.1702
14.4943 | -106.1767
14.498 | -106.1828
14.5017 | -106.1898
14.5055 | -106.2006
14.5092 | -106.2103
14.5129 | -106.2196
14.5167 | -106.2255
14.5204 | -106.2363
14.5241 | -106.2437
14.5278 | -106.2483
14.5316 | -106.25
14.5353 | -106.2519
14.539 | -106.252
14.5427 | -106.2504
14.5465 | -106.2502
14.5502 | -106.2503
14.5539 | -106.2462
14.5577 | -106.2412
14.5614 | -106.2347
14.5651 | -106.2275
14.5688 | -106.2193
14.5726 | -106.2054
14.5763 | -106.1934
14.58 | -106.1842
14.5837 | -106.1709
14.5875 | -106.1552
14.5912 | -106.1396
14.5949 | -106.1237
14.5986 | -106.1109
14.6024 | -106.0987
14.6061 | -106.0863
14.6098 | -106.0722
14.6136 | -106.0574
14.6173 | -106.0422
14.621 | -106.0296
14.6247 | -106.0199
14.6285 | -106.0098
14.6322 | -105.9992
14.6359 | -105.9855
14.6396 | -105.9766
14.6434 | -105.9714
14.6471 | -105.9679
14.6508 | -105.9658
14.6546 | -105.9645
14.6583 | -105.961
14.662 | -105.9599
14.6657 | -105.9628
14.6695 | -105.9662
14.6732 | -105.9703
14.6769 | -105.9749
14.6806 | -105.9812
14.6844 | -105.9891
14.6881 | -105.9964
14.6918 | -106.0034
14.6956 | -106.0111
14.6993 | -106.0216
14.703 | -106.0305
14.7067 | -106.0419
14.7105 | -106.057
14.7142 | -106.0675
14.7179 | -106.0823
14.7216 | -106.0976
14.7254 | -106.1101
14.7291 | -106.1203
14.7328 | -106.1285
14.7365 | -106.1379
14.7403 | -106.1485
14.744 | -106.1572
14.7477 | -106.1653
14.7515 | -106.1748
14.7552 | -106.183
14.7589 | -106.1885
14.7626 | -106.195
14.7664 | -106.2034
14.7701 | -106.2096
14.7738 | -106.2158
14.7775 | -106.2175
14.7813 | -106.2159
14.785 | -106.2182
14.7887 | -106.2155
14.7925 | -106.2114
14.7962 | -106.2109
14.7999 | -106.2099
14.8036 | -106.2068
14.8074 | -106.2016
14.8111 | -106.1912
14.8148 | -106.1831
14.8185 | -106.1736
14.8223 | -106.1694
14.826 | -106.1653
14.8297 | -106.158
14.8334 | -106.1547
14.8372 | -106.1511
14.8409 | -106.1458
14.8446 | -106.1421
14.8484 | -106.1399
14.8521 | -106.1359
14.8558 | -106.1323
14.8595 | -106.1283
14.8633 | -106.1294
14.867 | -106.1313
14.8707 | -106.1347
14.8744 | -106.1375
14.8782 | -106.1407
14.8819 | -106.1472
14.8856 | -106.1548
14.8894 | -106.1614
14.8931 | -106.168
14.8968 | -106.1755
14.9005 | -106.1859
14.9043 | -106.1949
14.908 | -106.2035
14.9117 | -106.2144
14.9154 | -106.2282
14.9192 | -106.2388
14.9229 | -106.2486
14.9266 | -106.2614
14.9304 | -106.272
14.9341 | -106.2823
14.9378 | -106.2934
14.9415 | -106.3073
14.9453 | -106.3188
14.949 | -106.3305
14.9527 | -106.3431
14.9564 | -106.3544
14.9602 | -106.3623
14.9639 | -106.3678
14.9676 | -106.3756
14.9713 | -106.3846
14.9751 | -106.3911
14.9788 | -106.3965
14.9825 | -106.3981
14.9863 | -106.3979
14.99 | -106.3992
14.9937 | -106.4036
14.9974 | -106.4038
15.0012 | -106.4059
15.0049 | -106.4036
15.0086 | -106.3975
15.0123 | -106.3907
15.0161 | -106.3843
15.0198 | -106.3776
15.0235 | -106.3717
15.0273 | -106.3598
15.031 | -106.3446
15.0347 | -106.3227
15.0384 | -106.3018
15.0422 | -106.2846
15.0459 | -106.2616
15.0496 | -106.2354
15.0533 | -106.2077
15.0571 | -106.1846
15.0608 | -106.1598
15.0645 | -106.1353
15.0683 | -106.1088
15.072 | -106.083
15.0757 | -106.0599
15.0794 | -106.0352
15.0832 | -106.0128
15.0869 | -105.9944
15.0906 | -105.9758
15.0943 | -105.9575
15.0981 | -105.9413
15.1018 | -105.9271
15.1055 | -105.919
15.1092 | -105.9143
15.113 | -105.9075
15.1167 | -105.9003
15.1204 | -105.8977
15.1242 | -105.899
15.1279 | -105.9005
15.1316 | -105.9046
15.1353 | -105.9096
15.1391 | -105.9138
15.1428 | -105.9199
15.1465 | -105.9266
15.1502 | -105.9351
15.154 | -105.9444
15.1577 | -105.9521
15.1614 | -105.9598
15.1652 | -105.9701
15.1689 | -105.9785
15.1726 | -105.9871
15.1763 | -105.9976
15.1801 | -106.0084
15.1838 | -106.0158
15.1875 | -106.0222
15.1912 | -106.0287
15.195 | -106.032
15.1987 | -106.0389
15.2024 | -106.0423
15.2061 | -106.0411
15.2099 | -106.0385
15.2136 | -106.0376
15.2173 | -106.0357
15.2211 | -106.0325
15.2248 | -106.0282
15.2285 | -106.0257
15.2322 | -106.0222
15.236 | -106.0169
15.2397 | -106.0108
15.2434 | -106.0047
15.2471 | -105.9992
15.2509 | -105.994
15.2546 | -105.9874
15.2583 | -105.9795
15.2621 | -105.972
15.2658 | -105.9649
15.2695 | -105.9563
15.2732 | -105.9474
15.277 | -105.9408
15.2807 | -105.9333
15.2844 | -105.9263
15.2881 | -105.9199
15.2919 | -105.9111
15.2956 | -105.9045
15.2993 | -105.9004
15.3031 | -105.8965
15.3068 | -105.8916
15.3105 | -105.8915
15.3142 | -105.8925
15.318 | -105.8921
15.3217 | -105.8952
15.3254 | -105.895
15.3291 | -105.8944
15.3329 | -105.8981
15.3366 | -105.903
15.3403 | -105.9073
15.344 | -105.9113
15.3478 | -105.9169
15.3515 | -105.9225
15.3552 | -105.9282
15.359 | -105.9349
15.3627 | -105.9393
15.3664 | -105.9432
15.3701 | -105.9481
15.3739 | -105.9514
15.3776 | -105.9547
15.3813 | -105.9574
15.385 | -105.9588
15.3888 | -105.9609
15.3925 | -105.963
15.3962 | -105.9672
15.4 | -105.9722
15.4037 | -105.9763
15.4074 | -105.9811
15.4111 | -105.9841
15.4149 | -105.9884
15.4186 | -105.9924
15.4223 | -105.9991
15.426 | -106.0066
15.4298 | -106.0113
15.4335 | -106.013
15.4372 | -106.0151
15.441 | -106.0214
15.4447 | -106.0291
15.4484 | -106.0331
15.4521 | -106.0371
15.4559 | -106.0414
15.4596 | -106.0483
15.4633 | -106.0553
15.467 | -106.0613
15.4708 | -106.066
15.4745 | -106.0678
15.4782 | -106.0695
15.4819 | -106.071
15.4857 | -106.0726
15.4894 | -106.0759
15.4931 | -106.0794
15.4969 | -106.08
15.5006 | -106.0798
15.5043 | -106.079
15.508 | -106.0803
15.5118 | -106.0833
15.5155 | -106.08
15.5192 | -106.0736
15.5229 | -106.0703
15.5267 | -106.0675
15.5304 | -106.0635
15.5341 | -106.0619
15.5379 | -106.0614
15.5416 | -106.0591
15.5453 | -106.0585
15.549 | -106.0562
15.5528 | -106.0513
15.5565 | -106.0481
15.5602 | -106.0446
15.5639 | -106.0405
15.5677 | -106.0368
15.5714 | -106.0337
15.5751 | -106.0313
15.5788 | -106.0274
15.5826 | -106.024
15.5863 | -106.0176
15.59 | -106.0109
15.5938 | -106.0054
15.5975 | -105.9999
15.6012 | -105.9957
15.6049 | -105.9917
15.6087 | -105.9843
15.6124 | -105.9754
15.6161 | -105.967
15.6198 | -105.9588
15.6236 | -105.951
15.6273 | -105.946
15.631 | -105.9404
15.6348 | -105.9357
15.6385 | -105.9323
15.6422 | -105.9278
15.6459 | -105.9223
15.6497 | -105.916
15.6534 | -105.9084
15.6571 | -105.9025
15.6608 | -105.8997
15.6646 | -105.8963
15.6683 | -105.8926
15.672 | -105.8886
15.6758 | -105.8851
15.6795 | -105.8829
15.6832 | -105.8813
15.6869 | -105.8781
15.6907 | -105.8756
15.6944 | -105.8758
15.6981 | -105.8731
15.7018 | -105.8704
15.7056 | -105.8704
15.7093 | -105.871
15.713 | -105.8713
15.7167 | -105.8716
15.7205 | -105.8734
15.7242 | -105.8776
15.7279 | -105.88
15.7317 | -105.8805
15.7354 | -105.8801
15.7391 | -105.883
15.7428 | -105.8848
15.7466 | -105.8865
15.7503 | -105.8882
15.754 | -105.8893
15.7577 | -105.8879
15.7615 | -105.887
15.7652 | -105.8859
15.7689 | -105.8844
15.7727 | -105.8805
15.7764 | -105.879
15.7801 | -105.8746
15.7838 | -105.8669
15.7876 | -105.8592
15.7913 | -105.8512
15.795 | -105.8456
15.7987 | -105.8399
15.8025 | -105.8356
15.8062 | -105.8318
15.8099 | -105.8293
15.8137 | -105.8265
15.8174 | -105.823
15.8211 | -105.8194
15.8248 | -105.8191
15.8286 | -105.8198
15.8323 | -105.8218
15.836 | -105.826
15.8397 | -105.8318
15.8435 | -105.8375
15.8472 | -105.8423
15.8509 | -105.8483
15.8546 | -105.8558
15.8584 | -105.8626
15.8621 | -105.8685
15.8658 | -105.8758
15.8696 | -105.8844
15.8733 | -105.8961
15.877 | -105.9064
15.8807 | -105.9183
15.8845 | -105.9294
15.8882 | -105.9377
15.8919 | -105.9447
15.8956 | -105.9553
15.8994 | -105.9667
15.9031 | -105.9774
15.9068 | -105.9862
15.9106 | -105.9911
15.9143 | -105.9987
15.918 | -106.0066
15.9217 | -106.0131
15.9255 | -106.0221
15.9292 | -106.0294
15.9329 | -106.0343
15.9366 | -106.0374
15.9404 | -106.0383
15.9441 | -106.0393
15.9478 | -106.0404
15.9515 | -106.0413
15.9553 | -106.0437
15.959 | -106.043
15.9627 | -106.0412
15.9665 | -106.0403
15.9702 | -106.0412
15.9739 | -106.0401
15.9776 | -106.0393
15.9814 | -106.04
15.9851 | -106.04
15.9888 | -106.0366
15.9925 | -106.0352
15.9963 | -106.0347
16 | -106.0368
16.0037 | -106.0385
16.0075 | -106.0406
16.0112 | -106.0438
16.0149 | -106.0449
16.0186 | -106.0472
16.0224 | -106.0483
16.0261 | -106.0511
16.0298 | -106.0557
16.0335 | -106.0597
16.0373 | -106.0636
16.041 | -106.068
16.0447 | -106.0729
16.0485 | -106.0778
16.0522 | -106.0839
16.0559 | -106.0899
16.0596 | -106.0954
16.0634 | -106.1013
16.0671 | -106.1058
16.0708 | -106.1107
16.0745 | -106.1152
16.0783 | -106.1184
16.082 | -106.1181
16.0857 | -106.1191
16.0894 | -106.1223
16.0932 | -106.1239
16.0969 | -106.1223
16.1006 | -106.1184
16.1044 | -106.1156
16.1081 | -106.1134
16.1118 | -106.1115
16.1155 | -106.1092
16.1193 | -106.1099
16.123 | -106.1068
16.1267 | -106.1064
16.1304 | -106.1082
16.1342 | -106.1078
16.1379 | -106.1073
16.1416 | -106.1087
16.1454 | -106.1079
16.1491 | -106.1065
16.1528 | -106.1038
16.1565 | -106.102
16.1603 | -106.1018
16.164 | -106.1001
16.1677 | -106.1001
16.1714 | -106.1006
16.1752 | -106.1013
16.1789 | -106.1016
16.1826 | -106.101
16.1864 | -106.0991
16.1901 | -106.1005
16.1938 | -106.1029
16.1975 | -106.1047
16.2013 | -106.1069
16.205 | -106.1101
16.2087 | -106.1151
16.2124 | -106.1197
16.2162 | -106.1261
16.2199 | -106.1365
16.2236 | -106.1448
16.2273 | -106.1529
16.2311 | -106.1634
16.2348 | -106.1726
16.2385 | -106.1815
16.2423 | -106.1912
16.246 | -106.2033
16.2497 | -106.2132
16.2534 | -106.2234
16.2572 | -106.2357
16.2609 | -106.2468
16.2646 | -106.2592
16.2683 | -106.2705
16.2721 | -106.2796
16.2758 | -106.2907
16.2795 | -106.3022
16.2833 | -106.3106
16.287 | -106.3207
16.2907 | -106.3322
16.2944 | -106.3408
16.2982 | -106.3485
16.3019 | -106.357
16.3056 | -106.3639
16.3093 | -106.3717
16.3131 | -106.3789
16.3168 | -106.3851
16.3205 | -106.3912
16.3242 | -106.3945
16.328 | -106.4005
16.3317 | -106.4062
16.3354 | -106.4122
16.3392 | -106.4143
16.3429 | -106.4158
16.3466 | -106.4185
16.3503 | -106.4194
16.3541 | -106.4198
16.3578 | -106.4216
16.3615 | -106.4203
16.3652 | -106.4212
16.369 | -106.425
16.3727 | -106.4292
16.3764 | -106.4309
16.3802 | -106.4308
16.3839 | -106.4326
16.3876 | -106.4348
16.3913 | -106.4356
16.3951 | -106.4353
16.3988 | -106.4358
16.4025 | -106.434
16.4062 | -106.4309
16.41 | -106.4283
16.4137 | -106.429
16.4174 | -106.4298
16.4212 | -106.428
16.4249 | -106.4251
16.4286 | -106.4224
16.4323 | -106.4181
16.4361 | -106.413
16.4398 | -106.4083
16.4435 | -106.4037
16.4472 | -106.4
16.451 | -106.3944
16.4547 | -106.3847
16.4584 | -106.3778
16.4621 | -106.3719
16.4659 | -106.3617
16.4696 | -106.3502
16.4733 | -106.3372
16.4771 | -106.3213
16.4808 | -106.3065
16.4845 | -106.295
16.4882 | -106.2833
16.492 | -106.2694
16.4957 | -106.2556
16.4994 | -106.2413
16.5031 | -106.2278
16.5069 | -106.2149
16.5106 | -106.1991
16.5143 | -106.1823
16.5181 | -106.1659
16.5218 | -106.1492
16.5255 | -106.1348
16.5292 | -106.124
16.533 | -106.1151
16.5367 | -106.1035
16.5404 | -106.0933
16.5441 | -106.0832
16.5479 | -106.0771
16.5516 | -106.0707
16.5553 | -106.0657
16.559 | -106.0631
16.5628 | -106.0612
16.5665 | -106.064
16.5702 | -106.0691
16.574 | -106.0764
16.5777 | -106.0861
16.5814 | -106.0947
16.5851 | -106.1063
16.5889 | -106.1218
16.5926 | -106.1375
16.5963 | -106.1584
16.6 | -106.18
16.6038 | -106.2013
16.6075 | -106.2223
16.6112 | -106.2463
16.615 | -106.2706
16.6187 | -106.2934
16.6224 | -106.3177
16.6261 | -106.3407
16.6299 | -106.3633
16.6336 | -106.3852
16.6373 | -106.406
16.641 | -106.4261
16.6448 | -106.4448
16.6485 | -106.4605
16.6522 | -106.4753
16.656 | -106.4911
16.6597 | -106.5048
16.6634 | -106.5159
16.6671 | -106.5189
16.6709 | -106.5283
16.6746 | -106.5356
16.6783 | -106.5411
16.682 | -106.5446
16.6858 | -106.5457
16.6895 | -106.549
16.6932 | -106.5528
16.6969 | -106.5532
16.7007 | -106.5522
16.7044 | -106.5517
16.7081 | -106.5526
16.7119 | -106.5565
16.7156 | -106.5585
16.7193 | -106.5537
16.723 | -106.5467
16.7268 | -106.5436
16.7305 | -106.5412
16.7342 | -106.5392
16.7379 | -106.5378
16.7417 | -106.539
16.7454 | -106.5412
16.7491 | -106.5432
16.7529 | -106.5421
16.7566 | -106.5453
16.7603 | -106.5464
16.764 | -106.546
16.7678 | -106.5455
16.7715 | -106.5466
16.7752 | -106.548
16.7789 | -106.5506
16.7827 | -106.5529
16.7864 | -106.5515
16.7901 | -106.5489
16.7938 | -106.5443
16.7976 | -106.5419
16.8013 | -106.5396
16.805 | -106.5355
16.8088 | -106.5315
16.8125 | -106.5274
16.8162 | -106.5236
16.8199 | -106.5207
16.8237 | -106.5159
16.8274 | -106.5119
16.8311 | -106.5083
16.8348 | -106.5037
16.8386 | -106.4994
16.8423 | -106.4947
16.846 | -106.493
16.8498 | -106.4916
16.8535 | -106.494
16.8572 | -106.4965
16.8609 | -106.4974
16.8647 | -106.4994
16.8684 | -106.501
16.8721 | -106.5047
16.8758 | -106.5105
16.8796 | -106.5175
16.8833 | -106.5232
16.887 | -106.5302
16.8908 | -106.5383
16.8945 | -106.5467
16.8982 | -106.5558
16.9019 | -106.567
16.9057 | -106.5764
16.9094 | -106.5874
16.9131 | -106.6
16.9168 | -106.6138
16.9206 | -106.6245
16.9243 | -106.6355
16.928 | -106.6475
16.9317 | -106.659
16.9355 | -106.6674
16.9392 | -106.6762
16.9429 | -106.6862
16.9467 | -106.6957
16.9504 | -106.7038
16.9541 | -106.7093
16.9578 | -106.7154
16.9616 | -106.7211
16.9653 | -106.725
16.969 | -106.7266
16.9727 | -106.7303
16.9765 | -106.7336
16.9802 | -106.7343
16.9839 | -106.7346
16.9877 | -106.7348
16.9914 | -106.7359
16.9951 | -106.741
16.9988 | -106.7417
17.0026 | -106.7427
17.0063 | -106.7431
17.01 | -106.7444
17.0137 | -106.7456
17.0175 | -106.7478
17.0212 | -106.746
17.0249 | -106.7431
17.0287 | -106.743
17.0324 | -106.7464
17.0361 | -106.7504
17.0398 | -106.7512
17.0436 | -106.7498
17.0473 | -106.7497
17.051 | -106.7527
17.0547 | -106.756
17.0585 | -106.7588
17.0622 | -106.7632
17.0659 | -106.7672
17.0696 | -106.7692
17.0734 | -106.7718
17.0771 | -106.7743
17.0808 | -106.7767
17.0846 | -106.7815
17.0883 | -106.7836
17.092 | -106.7821
17.0957 | -106.7814
17.0995 | -106.7803
17.1032 | -106.778
17.1069 | -106.7773
17.1106 | -106.776
17.1144 | -106.7746
17.1181 | -106.773
17.1218 | -106.7693
17.1256 | -106.7678
17.1293 | -106.7632
17.133 | -106.7577
17.1367 | -106.7551
17.1405 | -106.7528
17.1442 | -106.7503
17.1479 | -106.7458
17.1516 | -106.7428
17.1554 | -106.7395
17.1591 | -106.7358
17.1628 | -106.734
17.1666 | -106.7321
17.1703 | -106.7315
17.174 | -106.7308
17.1777 | -106.7328
17.1815 | -106.7334
17.1852 | -106.7298
17.1889 | -106.7278
17.1926 | -106.728
17.1964 | -106.7302
17.2001 | -106.7328
17.2038 | -106.7404
17.2075 | -106.7446
17.2113 | -106.7469
17.215 | -106.7473
17.2187 | -106.7479
17.2225 | -106.7496
17.2262 | -106.7537
17.2299 | -106.7581
17.2336 | -106.7616
17.2374 | -106.7648
17.2411 | -106.7685
17.2448 | -106.7741
17.2485 | -106.7789
17.2523 | -106.7844
17.256 | -106.7883
17.2597 | -106.7915
17.2635 | -106.7968
17.2672 | -106.8041
17.2709 | -106.8109
17.2746 | -106.8159
17.2784 | -106.8183
17.2821 | -106.8216
17.2858 | -106.8256
17.2895 | -106.8302
17.2933 | -106.8334
17.297 | -106.8336
17.3007 | -106.8341
17.3044 | -106.8344
17.3082 | -106.8338
17.3119 | -106.8344
17.3156 | -106.8358
17.3194 | -106.8345
17.3231 | -106.8342
17.3268 | -106.8358
17.3305 | -106.8368
17.3343 | -106.839
17.338 | -106.8432
17.3417 | -106.8463
17.3454 | -106.8479
17.3492 | -106.8522
17.3529 | -106.8542
17.3566 | -106.8569
17.3604 | -106.8582
17.3641 | -106.858
17.3678 | -106.8582
17.3715 | -106.8602
17.3753 | -106.8643
17.379 | -106.8681
17.3827 | -106.8691
17.3864 | -106.8679
17.3902 | -106.8675
17.3939 | -106.8722
17.3976 | -106.8719
17.4014 | -106.8711
17.4051 | -106.8701
17.4088 | -106.8673
17.4125 | -106.8638
17.4163 | -106.8594
17.42 | -106.8536
17.4237 | -106.8468
17.4274 | -106.8388
17.4312 | -106.8318
17.4349 | -106.8248
17.4386 | -106.8141
17.4423 | -106.8017
17.4461 | -106.7903
17.4498 | -106.7777
17.4535 | -106.7643
17.4573 | -106.7532
17.461 | -106.7428
17.4647 | -106.7342
17.4684 | -106.7252
17.4722 | -106.7166
17.4759 | -106.7086
17.4796 | -106.6993
17.4833 | -106.6902
17.4871 | -106.6824
17.4908 | -106.6737
17.4945 | -106.6646
17.4983 | -106.6611
17.502 | -106.6591
17.5057 | -106.6568
17.5094 | -106.6569
17.5132 | -106.6585
17.5169 | -106.6609
17.5206 | -106.6655
17.5243 | -106.6706
17.5281 | -106.6767
17.5318 | -106.6826
17.5355 | -106.6888
17.5392 | -106.6963
17.543 | -106.7053
17.5467 | -106.7162
17.5504 | -106.7271
17.5542 | -106.7371
17.5579 | -106.7468
17.5616 | -106.7557
17.5653 | -106.765
17.5691 | -106.7738
17.5728 | -106.7834
17.5765 | -106.7915
17.5802 | -106.7992
17.584 | -106.8077
17.5877 | -106.816
17.5914 | -106.8219
17.5952 | -106.8287
17.5989 | -106.8364
17.6026 | -106.844
17.6063 | -106.8506
17.6101 | -106.8573
17.6138 | -106.8615
17.6175 | -106.8652
17.6212 | -106.8656
17.625 | -106.8677
17.6287 | -106.8733
17.6324 | -106.8773
17.6362 | -106.8819
17.6399 | -106.8842
17.6436 | -106.8867
17.6473 | -106.8913
17.6511 | -106.8952
17.6548 | -106.8991
17.6585 | -106.9033
17.6622 | -106.9047
17.666 | -106.9046
17.6697 | -106.9045
17.6734 | -106.9044
17.6771 | -106.9057
17.6809 | -106.9059
17.6846 | -106.9053
17.6883 | -106.9068
17.6921 | -106.908
17.6958 | -106.9091
17.6995 | -106.9113
17.7032 | -106.909
17.707 | -106.9022
17.7107 | -106.8996
17.7144 | -106.8962
17.7181 | -106.8929
17.7219 | -106.8905
17.7256 | -106.8836
17.7293 | -106.8778
17.7331 | -106.8752
17.7368 | -106.8739
17.7405 | -106.8687
17.7442 | -106.8643
17.748 | -106.8581
17.7517 | -106.8517
17.7554 | -106.8451
17.7591 | -106.8415
17.7629 | -106.8364
17.7666 | -106.8294
17.7703 | -106.8235
17.7741 | -106.8167
17.7778 | -106.8093
17.7815 | -106.8041
17.7852 | -106.7978
17.789 | -106.7901
17.7927 | -106.7832
17.7964 | -106.7747
17.8001 | -106.7666
17.8039 | -106.7587
17.8076 | -106.7493
17.8113 | -106.7424
17.815 | -106.7366
17.8188 | -106.7309
17.8225 | -106.7271
17.8262 | -106.7224
17.83 | -106.7164
17.8337 | -106.7095
17.8374 | -106.7048
17.8411 | -106.7015
17.8449 | -106.7004
17.8486 | -106.7004
17.8523 | -106.6987
17.856 | -106.6998
17.8598 | -106.7012
17.8635 | -106.7026
17.8672 | -106.7063
17.871 | -106.7117
17.8747 | -106.7165
17.8784 | -106.7211
17.8821 | -106.7239
17.8859 | -106.7287
17.8896 | -106.7342
17.8933 | -106.7396
17.897 | -106.7445
17.9008 | -106.7485
17.9045 | -106.7536
17.9082 | -106.759
17.912 | -106.7632
17.9157 | -106.7676
17.9194 | -106.7716
17.9231 | -106.7758
17.9269 | -106.7796
17.9306 | -106.781
17.9343 | -106.783
17.938 | -106.7841
17.9418 | -106.7838
17.9455 | -106.7849
17.9492 | -106.7841
17.9529 | -106.7818
17.9567 | -106.7845
17.9604 | -106.7887
17.9641 | -106.7885
17.9679 | -106.7855
17.9716 | -106.7837
17.9753 | -106.7813
17.979 | -106.7819
17.9828 | -106.7808
17.9865 | -106.7759
17.9902 | -106.7689
17.9939 | -106.7645
17.9977 | -106.7629
18.0014 | -106.7632
18.0051 | -106.7644
18.0089 | -106.7608
18.0126 | -106.7586
18.0163 | -106.756
18.02 | -106.7547
18.0238 | -106.757
18.0275 | -106.7591
18.0312 | -106.7596
18.0349 | -106.7623
18.0387 | -106.7652
18.0424 | -106.7702
18.0461 | -106.7767
18.0498 | -106.7797
18.0536 | -106.7842
18.0573 | -106.7901
18.061 | -106.7967
18.0648 | -106.8053
18.0685 | -106.8107
18.0722 | -106.8141
18.0759 | -106.8172
18.0797 | -106.8208
18.0834 | -106.8248
18.0871 | -106.8302
18.0908 | -106.8351
18.0946 | -106.8386
18.0983 | -106.8398
18.102 | -106.8389
18.1058 | -106.8378
18.1095 | -106.8378
18.1132 | -106.8361
18.1169 | -106.8336
18.1207 | -106.8306
18.1244 | -106.8278
18.1281 | -106.8246
18.1318 | -106.8214
18.1356 | -106.8168
18.1393 | -106.812
18.143 | -106.807
18.1468 | -106.8016
18.1505 | -106.7965
18.1542 | -106.7922
18.1579 | -106.7907
18.1617 | -106.7887
18.1654 | -106.7863
18.1691 | -106.7838
18.1728 | -106.7826
18.1766 | -106.7822
18.1803 | -106.7848
18.184 | -106.7864
18.1877 | -106.7868
18.1915 | -106.7879
18.1952 | -106.79
18.1989 | -106.7924
18.2027 | -106.7957
18.2064 | -106.8003
18.2101 | -106.8039
18.2138 | -106.8059
18.2176 | -106.8093
18.2213 | -106.8125
18.225 | -106.8175
18.2287 | -106.8193
18.2325 | -106.8209
18.2362 | -106.8224
18.2399 | -106.8242
18.2437 | -106.8268
18.2474 | -106.8285
18.2511 | -106.8264
18.2548 | -106.8196
18.2586 | -106.8118
18.2623 | -106.8064
18.266 | -106.8018
18.2697 | -106.7943
18.2735 | -106.787
18.2772 | -106.779
18.2809 | -106.7669
18.2846 | -106.7576
18.2884 | -106.7484
18.2921 | -106.7408
18.2958 | -106.7327
18.2996 | -106.7226
18.3033 | -106.7139
18.307 | -106.7058
18.3107 | -106.6981
18.3145 | -106.6903
18.3182 | -106.6822
18.3219 | -106.675
18.3256 | -106.6685
18.3294 | -106.6628
18.3331 | -106.6587
18.3368 | -106.6561
18.3406 | -106.6527
18.3443 | -106.6517
18.348 | -106.6517
18.3517 | -106.6544
18.3555 | -106.6571
18.3592 | -106.6616
18.3629 | -106.6674
18.3666 | -106.674
18.3704 | -106.6809
18.3741 | -106.6891
18.3778 | -106.6959
18.3816 | -106.7029
18.3853 | -106.7112
18.389 | -106.717
18.3927 | -106.7209
18.3965 | -106.7214
18.4002 | -106.7233
18.4039 | -106.7273
18.4076 | -106.7305
18.4114 | -106.7333
18.4151 | -106.7343
18.4188 | -106.7333
18.4225 | -106.731
18.4263 | -106.7281
18.43 | -106.725
18.4337 | -106.7193
18.4375 | -106.7086
18.4412 | -106.6942
18.4449 | -106.6806
18.4486 | -106.6651
18.4524 | -106.6466
18.4561 | -106.626
18.4598 | -106.6065
18.4635 | -106.5858
18.4673 | -106.5639
18.471 | -106.542
18.4747 | -106.5168
18.4785 | -106.4908
18.4822 | -106.4644
18.4859 | -106.4384
18.4896 | -106.4142
18.4934 | -106.3901
18.4971 | -106.3672
18.5008 | -106.3455
18.5045 | -106.3248
18.5083 | -106.3065
18.512 | -106.2911
18.5157 | -106.2788
18.5194 | -106.268
18.5232 | -106.258
18.5269 | -106.2502
18.5306 | -106.245
18.5344 | -106.2425
18.5381 | -106.2407
18.5418 | -106.2411
18.5455 | -106.2445
18.5493 | -106.2503
18.553 | -106.259
18.5567 | -106.2706
18.5604 | -106.2839
18.5642 | -106.2995
18.5679 | -106.3142
18.5716 | -106.3321
18.5754 | -106.3534
18.5791 | -106.3758
18.5828 | -106.3986
18.5865 | -106.4212
18.5903 | -106.4457
18.594 | -106.4711
18.5977 | -106.4994
18.6014 | -106.5273
18.6052 | -106.5564
18.6089 | -106.584
18.6126 | -106.6113
18.6164 | -106.6381
18.6201 | -106.6641
18.6238 | -106.6906
18.6275 | -106.7197
18.6313 | -106.7473
18.635 | -106.7736
18.6387 | -106.8004
18.6424 | -106.8305
18.6462 | -106.8598
18.6499 | -106.8845
18.6536 | -106.9078
18.6573 | -106.9324
18.6611 | -106.958
18.6648 | -106.9783
18.6685 | -106.9935
18.6723 | -107.0087
18.676 | -107.0216
18.6797 | -107.0326
18.6834 | -107.0413
18.6872 | -107.0466
18.6909 | -107.0484
18.6946 | -107.0513
18.6983 | -107.0526
18.7021 | -107.0534
18.7058 | -107.0511
18.7095 | -107.0468
18.7133 | -107.0419
18.717 | -107.0352
18.7207 | -107.0273
18.7244 | -107.0158
18.7282 | -107.0012
18.7319 | -106.9873
18.7356 | -106.9731
18.7393 | -106.9616
18.7431 | -106.9494
18.7468 | -106.9364
18.7505 | -106.9226
18.7543 | -106.9075
18.758 | -106.8947
18.7617 | -106.8819
18.7654 | -106.8672
18.7692 | -106.8512
18.7729 | -106.8372
18.7766 | -106.825
18.7803 | -106.8132
18.7841 | -106.8018
18.7878 | -106.791
18.7915 | -106.782
18.7952 | -106.7747
18.799 | -106.7663
18.8027 | -106.7604
18.8064 | -106.7547
18.8102 | -106.7503
18.8139 | -106.7471
18.8176 | -106.7478
18.8213 | -106.7461
18.8251 | -106.7453
18.8288 | -106.7448
18.8325 | -106.7448
18.8362 | -106.7462
18.84 | -106.7454
18.8437 | -106.7459
18.8474 | -106.7458
18.8512 | -106.7481
18.8549 | -106.752
18.8586 | -106.7543
18.8623 | -106.7589
18.8661 | -106.7615
18.8698 | -106.7646
18.8735 | -106.7694
18.8772 | -106.7749
18.881 | -106.7789
18.8847 | -106.7847
18.8884 | -106.7924
18.8922 | -106.7983
18.8959 | -106.8022
18.8996 | -106.8063
18.9033 | -106.8109
18.9071 | -106.8134
18.9108 | -106.8157
18.9145 | -106.8156
18.9182 | -106.8178
18.922 | -106.8217
18.9257 | -106.8252
18.9294 | -106.8286
18.9331 | -106.8318
18.9369 | -106.8349
18.9406 | -106.8376
18.9443 | -106.8396
18.9481 | -106.8423
18.9518 | -106.8447
18.9555 | -106.8489
18.9592 | -106.8528
18.963 | -106.8522
18.9667 | -106.8535
18.9704 | -106.8552
18.9741 | -106.8554
18.9779 | -106.8559
18.9816 | -106.8595
18.9853 | -106.8633
18.9891 | -106.8656
18.9928 | -106.8669
18.9965 | -106.8706
19.0002 | -106.8725
19.004 | -106.8734
19.0077 | -106.8743
19.0114 | -106.8766
19.0151 | -106.8798
19.0189 | -106.8833
19.0226 | -106.8849
19.0263 | -106.8854
19.03 | -106.8861
19.0338 | -106.8867
19.0375 | -106.8914
19.0412 | -106.8929
19.045 | -106.8924
19.0487 | -106.895
19.0524 | -106.8971
19.0561 | -106.8967
19.0599 | -106.8965
19.0636 | -106.8954
19.0673 | -106.8947
19.071 | -106.8945
19.0748 | -106.8906
19.0785 | -106.8853
19.0822 | -106.8821
19.086 | -106.8799
19.0897 | -106.8766
19.0934 | -106.8715
19.0971 | -106.8642
19.1009 | -106.857
19.1046 | -106.85
19.1083 | -106.8421
19.112 | -106.8349
19.1158 | -106.8264
19.1195 | -106.8179
19.1232 | -106.8058
19.127 | -106.7935
19.1307 | -106.7799
19.1344 | -106.7624
19.1381 | -106.7447
19.1419 | -106.7271
19.1456 | -106.7103
19.1493 | -106.6939
19.153 | -106.6778
19.1568 | -106.6596
19.1605 | -106.6414
19.1642 | -106.625
19.1679 | -106.6073
19.1717 | -106.5903
19.1754 | -106.5751
19.1791 | -106.5621
19.1829 | -106.5487
19.1866 | -106.5367
19.1903 | -106.5238
19.194 | -106.5119
19.1978 | -106.5022
19.2015 | -106.4926
19.2052 | -106.4841
19.2089 | -106.4764
19.2127 | -106.4684
19.2164 | -106.4619
19.2201 | -106.4573
19.2239 | -106.4541
19.2276 | -106.4502
19.2313 | -106.4497
19.235 | -106.4524
19.2388 | -106.4528
19.2425 | -106.453
19.2462 | -106.4558
19.2499 | -106.4556
19.2537 | -106.453
19.2574 | -106.4529
19.2611 | -106.4541
19.2648 | -106.4563
19.2686 | -106.4559
19.2723 | -106.457
19.276 | -106.4585
19.2798 | -106.4586
19.2835 | -106.4595
19.2872 | -106.4596
19.2909 | -106.4592
19.2947 | -106.4619
19.2984 | -106.4628
19.3021 | -106.4618
19.3058 | -106.4632
19.3096 | -106.4603
19.3133 | -106.4572
19.317 | -106.4544
19.3208 | -106.453
19.3245 | -106.4526
19.3282 | -106.4519
19.3319 | -106.4509
19.3357 | -106.4498
19.3394 | -106.447
19.3431 | -106.4423
19.3468 | -106.4426
19.3506 | -106.4408
19.3543 | -106.4364
19.358 | -106.4293
19.3618 | -106.4225
19.3655 | -106.4165
19.3692 | -106.412
19.3729 | -106.4108
19.3767 | -106.41
19.3804 | -106.4061
19.3841 | -106.4033
19.3878 | -106.4024
19.3916 | -106.3994
19.3953 | -106.3971
19.399 | -106.3952
19.4027 | -106.3941
19.4065 | -106.395
19.4102 | -106.396
19.4139 | -106.3959
19.4177 | -106.3954
19.4214 | -106.3923
19.4251 | -106.3891
19.4288 | -106.3859
19.4326 | -106.3825
19.4363 | -106.3784
19.44 | -106.3737
19.4437 | -106.371
19.4475 | -106.3673
19.4512 | -106.362
19.4549 | -106.3566
19.4587 | -106.3534
19.4624 | -106.3508
19.4661 | -106.3479
19.4698 | -106.3425
19.4736 | -106.3353
19.4773 | -106.3294
19.481 | -106.3229
19.4847 | -106.3168
19.4885 | -106.311
19.4922 | -106.3045
19.4959 | -106.2977
19.4997 | -106.2932
19.5034 | -106.2898
19.5071 | -106.285
19.5108 | -106.2795
19.5146 | -106.2732
19.5183 | -106.2676
19.522 | -106.2627
19.5257 | -106.2604
19.5295 | -106.2599
19.5332 | -106.2588
19.5369 | -106.2553
19.5406 | -106.2506
19.5444 | -106.2493
19.5481 | -106.2496
19.5518 | -106.2489
19.5556 | -106.2491
19.5593 | -106.2485
19.563 | -106.2492
19.5667 | -106.2507
19.5705 | -106.2511
19.5742 | -106.2528
19.5779 | -106.2546
19.5816 | -106.2563
19.5854 | -106.2593
19.5891 | -106.262
19.5928 | -106.2651
19.5966 | -106.2655
19.6003 | -106.2663
19.604 | -106.268
19.6077 | -106.2693
19.6115 | -106.2717
19.6152 | -106.2756
19.6189 | -106.2784
19.6226 | -106.2788
19.6264 | -106.2774
19.6301 | -106.2759
19.6338 | -106.2781
19.6375 | -106.2814
19.6413 | -106.2832
19.645 | -106.2815
19.6487 | -106.2797
19.6525 | -106.2774
19.6562 | -106.276
19.6599 | -106.2775
19.6636 | -106.2782
19.6674 | -106.278
19.6711 | -106.2777
19.6748 | -106.2767
19.6785 | -106.2765
19.6823 | -106.2753
19.686 | -106.2742
19.6897 | -106.274
19.6935 | -106.2727
19.6972 | -106.2687
19.7009 | -106.266
19.7046 | -106.2644
19.7084 | -106.2627
19.7121 | -106.2595
19.7158 | -106.2565
19.7195 | -106.2562
19.7233 | -106.2526
19.727 | -106.2489
19.7307 | -106.2467
19.7345 | -106.2438
19.7382 | -106.2393
19.7419 | -106.2349
19.7456 | -106.2312
19.7494 | -106.2284
19.7531 | -106.2259
19.7568 | -106.2227
19.7605 | -106.2201
19.7643 | -106.2163
19.768 | -106.2126
19.7717 | -106.2094
19.7754 | -106.2066
19.7792 | -106.2048
19.7829 | -106.2006
19.7866 | -106.1967
19.7904 | -106.1936
19.7941 | -106.1945
19.7978 | -106.1949
19.8015 | -106.1939
19.8053 | -106.1953
19.809 | -106.1956
19.8127 | -106.1944
19.8164 | -106.1972
19.8202 | -106.2007
19.8239 | -106.2021
19.8276 | -106.2033
19.8314 | -106.2052
19.8351 | -106.2061
19.8388 | -106.2089
19.8425 | -106.2123
19.8463 | -106.2155
19.85 | -106.2213
19.8537 | -106.2268
19.8574 | -106.2313
19.8612 | -106.234
19.8649 | -106.2392
19.8686 | -106.2462
19.8724 | -106.2511
19.8761 | -106.2566
19.8798 | -106.2623
19.8835 | -106.2678
19.8873 | -106.2732
19.891 | -106.2807
19.8947 | -106.286
19.8984 | -106.2913
19.9022 | -106.2977
19.9059 | -106.3033
19.9096 | -106.3101
19.9133 | -106.319
19.9171 | -106.3249
19.9208 | -106.3287
19.9245 | -106.3314
19.9283 | -106.3366
19.932 | -106.3416
19.9357 | -106.3464
19.9394 | -106.3504
19.9432 | -106.3531
19.9469 | -106.358
19.9506 | -106.3616
19.9543 | -106.3653
19.9581 | -106.3706
19.9618 | -106.3769
19.9655 | -106.3839
19.9693 | -106.3897
19.973 | -106.3951
19.9767 | -106.4028
19.9804 | -106.4112
19.9842 | -106.4184
19.9879 | -106.4239
19.9916 | -106.4291
19.9953 | -106.4342
19.9991 | -106.4382
20.0028 | -106.4402
20.0065 | -106.4444
20.0102 | -106.4479
20.014 | -106.4488
20.0177 | -106.4502
20.0214 | -106.452
20.0252 | -106.4544
20.0289 | -106.4563
20.0326 | -106.4581
20.0363 | -106.459
20.0401 | -106.4605
20.0438 | -106.4632
20.0475 | -106.465
20.0512 | -106.4673
20.055 | -106.471
20.0587 | -106.4725
20.0624 | -106.4731
20.0662 | -106.4766
20.0699 | -106.4788
20.0736 | -106.4807
20.0773 | -106.4818
20.0811 | -106.4836
20.0848 | -106.4875
20.0885 | -106.4901
20.0922 | -106.4922
20.096 | -106.4953
20.0997 | -106.4983
20.1034 | -106.5006
20.1072 | -106.5033
20.1109 | -106.5057
20.1146 | -106.5109
20.1183 | -106.516
20.1221 | -106.5182
20.1258 | -106.5188
20.1295 | -106.5216
20.1332 | -106.5229
20.137 | -106.5257
20.1407 | -106.53
20.1444 | -106.5336
20.1481 | -106.5337
20.1519 | -106.5329
20.1556 | -106.5348
20.1593 | -106.5355
20.1631 | -106.5352
20.1668 | -106.5367
20.1705 | -106.5382
20.1742 | -106.5379
20.178 | -106.5379
20.1817 | -106.5364
20.1854 | -106.5315
20.1891 | -106.5279
20.1929 | -106.5264
20.1966 | -106.5259
20.2003 | -106.5227
20.2041 | -106.5183
20.2078 | -106.5158
20.2115 | -106.5153
20.2152 | -106.5135
20.219 | -106.5101
20.2227 | -106.5063
20.2264 | -106.5003
20.2301 | -106.4944
20.2339 | -106.4868
20.2376 | -106.4793
20.2413 | -106.4709
20.245 | -106.4605
20.2488 | -106.4507
20.2525 | -106.4387
20.2562 | -106.4246
20.26 | -106.4121
20.2637 | -106.3986
20.2674 | -106.384
20.2711 | -106.3685
20.2749 | -106.3507
20.2786 | -106.3343
20.2823 | -106.3185
20.286 | -106.3004
20.2898 | -106.2808
20.2935 | -106.2613
20.2972 | -106.243
20.301 | -106.2262
20.3047 | -106.2081
20.3084 | -106.1886
20.3121 | -106.1684
20.3159 | -106.1492
20.3196 | -106.1321
20.3233 | -106.114
20.327 | -106.0963
20.3308 | -106.0796
20.3345 | -106.0639
20.3382 | -106.0465
20.342 | -106.0301
20.3457 | -106.0175
20.3494 | -106.0047
20.3531 | -105.9939
20.3569 | -105.9843
20.3606 | -105.9743
20.3643 | -105.9651
20.368 | -105.9591
20.3718 | -105.9532
20.3755 | -105.9488
20.3792 | -105.9448
20.3829 | -105.9411
20.3867 | -105.9377
20.3904 | -105.9361
20.3941 | -105.9347
20.3979 | -105.9348
20.4016 | -105.9348
20.4053 | -105.9358
20.409 | -105.9376
20.4128 | -105.9423
20.4165 | -105.946
20.4202 | -105.949
20.4239 | -105.95
20.4277 | -105.9525
20.4314 | -105.9537
20.4351 | -105.9572
20.4389 | -105.9625
20.4426 | -105.9669
20.4463 | -105.9716
20.45 | -105.977
20.4538 | -105.9815
20.4575 | -105.9856
20.4612 | -105.9895
20.4649 | -105.9938
20.4687 | -105.9983
20.4724 | -106.0048
20.4761 | -106.0127
20.4799 | -106.0189
20.4836 | -106.023
20.4873 | -106.0272
20.491 | -106.0314
20.4948 | -106.0368
20.4985 | -106.0411
20.5022 | -106.0449
20.5059 | -106.047
20.5097 | -106.0498
20.5134 | -106.0552
20.5171 | -106.0637
20.5208 | -106.0704
20.5246 | -106.0773
20.5283 | -106.0851
20.532 | -106.0917
20.5358 | -106.0984
20.5395 | -106.1049
20.5432 | -106.1139
20.5469 | -106.1222
20.5507 | -106.127
20.5544 | -106.1343
20.5581 | -106.1419
20.5618 | -106.15
20.5656 | -106.1577
20.5693 | -106.1643
20.573 | -106.17
20.5768 | -106.1778
20.5805 | -106.1858
20.5842 | -106.1914
20.5879 | -106.1954
20.5917 | -106.1994
20.5954 | -106.2032
20.5991 | -106.2067
20.6028 | -106.2087
20.6066 | -106.2118
20.6103 | -106.2127
20.614 | -106.2148
20.6178 | -106.2156
20.6215 | -106.2133
20.6252 | -106.2115
20.6289 | -106.2101
20.6327 | -106.2077
20.6364 | -106.2035
20.6401 | -106.2005
20.6438 | -106.1966
20.6476 | -106.1947
20.6513 | -106.1916
20.655 | -106.1882
20.6587 | -106.1831
20.6625 | -106.1784
20.6662 | -106.173
20.6699 | -106.1702
20.6737 | -106.1685
20.6774 | -106.166
20.6811 | -106.1618
20.6848 | -106.1585
20.6886 | -106.1549
20.6923 | -106.1515
20.696 | -106.1467
20.6997 | -106.1417
20.7035 | -106.1362
20.7072 | -106.13
20.7109 | -106.1255
20.7147 | -106.1223
20.7184 | -106.1185
20.7221 | -106.116
20.7258 | -106.1146
20.7296 | -106.1106
20.7333 | -106.1034
20.737 | -106.0995
20.7407 | -106.0946
20.7445 | -106.0898
20.7482 | -106.0853
20.7519 | -106.0796
20.7556 | -106.0748
20.7594 | -106.0682
20.7631 | -106.0604
20.7668 | -106.055
20.7706 | -106.0492
20.7743 | -106.0421
20.778 | -106.0361
20.7817 | -106.0283
20.7855 | -106.0178
20.7892 | -106.0083
20.7929 | -105.9987
20.7966 | -105.9887
20.8004 | -105.9805
20.8041 | -105.9707
20.8078 | -105.9565
20.8116 | -105.9434
20.8153 | -105.933
20.819 | -105.9216
20.8227 | -105.9082
20.8265 | -105.8988
20.8302 | -105.8904
20.8339 | -105.8793
20.8376 | -105.8662
20.8414 | -105.8572
20.8451 | -105.8507
20.8488 | -105.8422
20.8526 | -105.8342
20.8563 | -105.8259
20.86 | -105.8219
20.8637 | -105.8183
20.8675 | -105.8146
20.8712 | -105.811
20.8749 | -105.8051
20.8786 | -105.8036
20.8824 | -105.8036
20.8861 | -105.8048
20.8898 | -105.8067
20.8935 | -105.8096
20.8973 | -105.8155
20.901 | -105.8199
20.9047 | -105.8261
20.9085 | -105.8312
20.9122 | -105.8369
20.9159 | -105.8449
20.9196 | -105.8545
20.9234 | -105.8662
20.9271 | -105.8758
20.9308 | -105.8864
20.9345 | -105.8971
20.9383 | -105.9083
20.942 | -105.9191
20.9457 | -105.9299
20.9495 | -105.9402
20.9532 | -105.9504
20.9569 | -105.96
20.9606 | -105.9713
20.9644 | -105.9831
20.9681 | -105.9916
20.9718 | -106.0001
20.9755 | -106.0086
20.9793 | -106.0182
20.983 | -106.0277
20.9867 | -106.0353
20.9904 | -106.0414
20.9942 | -106.0494
20.9979 | -106.0559
21.0016 | -106.0615
21.0054 | -106.067
21.0091 | -106.0728
21.0128 | -106.0773
21.0165 | -106.0798
21.0203 | -106.0806
21.024 | -106.081
21.0277 | -106.0828
21.0314 | -106.0862
21.0352 | -106.0865
21.0389 | -106.0857
21.0426 | -106.0851
21.0464 | -106.0847
21.0501 | -106.0837
21.0538 | -106.0822
21.0575 | -106.0815
21.0613 | -106.0772
21.065 | -106.0729
21.0687 | -106.0681
21.0724 | -106.0633
21.0762 | -106.0584
21.0799 | -106.0532
21.0836 | -106.0496
21.0874 | -106.0456
21.0911 | -106.0414
21.0948 | -106.0352
21.0985 | -106.0294
21.1023 | -106.0255
21.106 | -106.0239
21.1097 | -106.0206
21.1134 | -106.0169
21.1172 | -106.0131
21.1209 | -106.0108
21.1246 | -106.0059
21.1283 | -105.9991
21.1321 | -105.9938
21.1358 | -105.9908
21.1395 | -105.9878
21.1433 | -105.9813
21.147 | -105.9762
21.1507 | -105.9737
21.1544 | -105.9724
21.1582 | -105.9669
21.1619 | -105.9617
21.1656 | -105.9575
21.1693 | -105.9528
21.1731 | -105.9523
21.1768 | -105.9503
21.1805 | -105.9486
21.1843 | -105.946
21.188 | -105.9445
21.1917 | -105.9433
21.1954 | -105.9412
21.1992 | -105.9391
21.2029 | -105.9384
21.2066 | -105.9402
21.2103 | -105.9425
21.2141 | -105.9444
21.2178 | -105.9461
21.2215 | -105.9472
21.2252 | -105.9502
21.229 | -105.9522
21.2327 | -105.9544
21.2364 | -105.9578
21.2402 | -105.9615
21.2439 | -105.9677
21.2476 | -105.9747
21.2513 | -105.9821
21.2551 | -105.9894
21.2588 | -105.9962
21.2625 | -106.0025
21.2662 | -106.009
21.27 | -106.0176
21.2737 | -106.0285
21.2774 | -106.0396
21.2812 | -106.0513
21.2849 | -106.0613
21.2886 | -106.0723
21.2923 | -106.0852
21.2961 | -106.097
21.2998 | -106.1093
21.3035 | -106.1214
21.3072 | -106.1353
21.311 | -106.1479
21.3147 | -106.1612
21.3184 | -106.1754
21.3222 | -106.1908
21.3259 | -106.2067
21.3296 | -106.2205
21.3333 | -106.2341
21.3371 | -106.2477
21.3408 | -106.2614
21.3445 | -106.2747
21.3482 | -106.2889
21.352 | -106.3043
21.3557 | -106.3187
21.3594 | -106.3345
21.3631 | -106.3513
21.3669 | -106.3637
21.3706 | -106.3762
21.3743 | -106.3873
21.3781 | -106.3959
21.3818 | -106.4038
21.3855 | -106.4119
21.3892 | -106.4187
21.393 | -106.4258
21.3967 | -106.4345
21.4004 | -106.4399
21.4041 | -106.4413
21.4079 | -106.4407
21.4116 | -106.4401
21.4153 | -106.4399
21.4191 | -106.4387
21.4228 | -106.4367
21.4265 | -106.4323
21.4302 | -106.4275
21.434 | -106.4218
21.4377 | -106.4143
21.4414 | -106.4051
21.4451 | -106.3981
21.4489 | -106.3881
21.4526 | -106.3732
21.4563 | -106.3577
21.4601 | -106.3441
21.4638 | -106.3289
21.4675 | -106.3129
21.4712 | -106.2981
21.475 | -106.2812
21.4787 | -106.2607
21.4824 | -106.2386
21.4861 | -106.2153
21.4899 | -106.1938
21.4936 | -106.1704
21.4973 | -106.1456
21.501 | -106.1219
21.5048 | -106.1007
21.5085 | -106.0782
21.5122 | -106.0562
21.516 | -106.034
21.5197 | -106.0133
21.5234 | -105.994
21.5271 | -105.9755
21.5309 | -105.9575
21.5346 | -105.9407
21.5383 | -105.9251
21.542 | -105.9098
21.5458 | -105.8958
21.5495 | -105.8828
21.5532 | -105.8725
21.557 | -105.8622
21.5607 | -105.8524
21.5644 | -105.845
21.5681 | -105.8394
21.5719 | -105.8354
21.5756 | -105.832
21.5793 | -105.8291
21.583 | -105.8274
21.5868 | -105.8259
21.5905 | -105.8265
21.5942 | -105.8293
21.598 | -105.8338
21.6017 | -105.8389
21.6054 | -105.8452
21.6091 | -105.8523
21.6129 | -105.8599
21.6166 | -105.8698
21.6203 | -105.8807
21.624 | -105.8916
21.6278 | -105.9013
21.6315 | -105.9109
21.6352 | -105.9232
21.6389 | -105.9366
21.6427 | -105.9506
21.6464 | -105.9665
21.6501 | -105.9823
21.6539 | -105.997
21.6576 | -106.0089
21.6613 | -106.0206
21.665 | -106.0328
21.6688 | -106.046
21.6725 | -106.0585
21.6762 | -106.0701
21.6799 | -106.0827
21.6837 | -106.0937
21.6874 | -106.1045
21.6911 | -106.1132
21.6949 | -106.1191
21.6986 | -106.1241
21.7023 | -106.1294
21.706 | -106.1334
21.7098 | -106.1373
21.7135 | -106.1383
21.7172 | -106.1378
21.7209 | -106.1379
21.7247 | -106.1371
21.7284 | -106.1363
21.7321 | -106.1363
21.7358 | -106.1351
21.7396 | -106.1331
21.7433 | -106.1317
21.747 | -106.1289
21.7508 | -106.1252
21.7545 | -106.122
21.7582 | -106.1212
21.7619 | -106.1197
21.7657 | -106.1173
21.7694 | -106.1142
21.7731 | -106.1095
21.7768 | -106.1026
21.7806 | -106.0975
21.7843 | -106.0956
21.788 | -106.0936
21.7918 | -106.092
21.7955 | -106.0889
21.7992 | -106.0877
21.8029 | -106.0848
21.8067 | -106.0829
21.8104 | -106.0791
21.8141 | -106.0773
21.8178 | -106.0759
21.8216 | -106.0727
21.8253 | -106.0706
21.829 | -106.07
21.8328 | -106.0687
21.8365 | -106.067
21.8402 | -106.0657
21.8439 | -106.0637
21.8477 | -106.0619
21.8514 | -106.0591
21.8551 | -106.055
21.8588 | -106.0511
21.8626 | -106.0467
21.8663 | -106.0436
21.87 | -106.0422
21.8737 | -106.0408
21.8775 | -106.0394
21.8812 | -106.0388
21.8849 | -106.0376
21.8887 | -106.0354
21.8924 | -106.0339
21.8961 | -106.0335
21.8998 | -106.0324
21.9036 | -106.0311
21.9073 | -106.0286
21.911 | -106.0262
21.9147 | -106.0242
21.9185 | -106.0233
21.9222 | -106.0224
21.9259 | -106.019
21.9297 | -106.0169
21.9334 | -106.0153
21.9371 | -106.0134
21.9408 | -106.0116
21.9446 | -106.0096
21.9483 | -106.0068
21.952 | -106.0036
21.9557 | -105.9985
21.9595 | -105.991
21.9632 | -105.9841
21.9669 | -105.9806
21.9706 | -105.9755
21.9744 | -105.9688
21.9781 | -105.9618
21.9818 | -105.9562
21.9856 | -105.9513
21.9893 | -105.9469
21.993 | -105.9419
21.9967 | -105.9359
22.0005 | -105.9293
22.0042 | -105.9237
22.0079 | -105.9163
22.0116 | -105.9072
22.0154 | -105.8989
22.0191 | -105.8896
22.0228 | -105.8825
22.0266 | -105.8768
22.0303 | -105.8665
22.034 | -105.857
22.0377 | -105.849
22.0415 | -105.8397
22.0452 | -105.8326
22.0489 | -105.8234
22.0526 | -105.816
22.0564 | -105.8081
22.0601 | -105.8006
22.0638 | -105.7942
22.0676 | -105.7881
22.0713 | -105.7831
22.075 | -105.7796
22.0787 | -105.7766
22.0825 | -105.7753
22.0862 | -105.7761
22.0899 | -105.7779
22.0936 | -105.7791
22.0974 | -105.7787
22.1011 | -105.7809
22.1048 | -105.7838
22.1085 | -105.787
22.1123 | -105.7921
22.116 | -105.7986
22.1197 | -105.8025
22.1235 | -105.8061
22.1272 | -105.8093
22.1309 | -105.8132
22.1346 | -105.8203
22.1384 | -105.8292
22.1421 | -105.836
22.1458 | -105.8419
22.1495 | -105.848
22.1533 | -105.8554
22.157 | -105.8637
22.1607 | -105.8713
22.1645 | -105.8789
22.1682 | -105.8863
22.1719 | -105.8952
22.1756 | -105.9022
22.1794 | -105.9072
22.1831 | -105.91
22.1868 | -105.9121
22.1905 | -105.9147
22.1943 | -105.9161
22.198 | -105.9193
22.2017 | -105.9228
22.2055 | -105.924
22.2092 | -105.925
22.2129 | -105.9254
22.2166 | -105.9245
22.2204 | -105.9239
22.2241 | -105.9234
22.2278 | -105.9232
22.2315 | -105.9218
22.2353 | -105.9193
22.239 | -105.9171
22.2427 | -105.914
22.2464 | -105.912
22.2502 | -105.9114
22.2539 | -105.9102
22.2576 | -105.9077
22.2614 | -105.9047
22.2651 | -105.9021
22.2688 | -105.9004
22.2725 | -105.9021
22.2763 | -105.9028
22.28 | -105.9022
22.2837 | -105.9032
22.2874 | -105.9047
22.2912 | -105.9061
22.2949 | -105.9079
22.2986 | -105.9083
22.3024 | -105.9065
22.3061 | -105.9071
22.3098 | -105.9095
22.3135 | -105.9133
22.3173 | -105.918
22.321 | -105.923
22.3247 | -105.9277
22.3284 | -105.9328
22.3322 | -105.9395
22.3359 | -105.9464
22.3396 | -105.9531
22.3434 | -105.9621
22.3471 | -105.9693
22.3508 | -105.975
22.3545 | -105.9838
22.3583 | -105.992
22.362 | -106.0002
22.3657 | -106.0082
22.3694 | -106.0145
22.3732 | -106.0218
22.3769 | -106.0296
22.3806 | -106.0355
22.3843 | -106.0436
22.3881 | -106.0512
22.3918 | -106.0578
22.3955 | -106.064
22.3993 | -106.0715
22.403 | -106.0782
22.4067 | -106.0843
22.4104 | -106.0895
22.4142 | -106.0961
22.4179 | -106.1024
22.4216 | -106.1098
22.4253 | -106.1144
22.4291 | -106.1168
22.4328 | -106.1215
22.4365 | -106.1271
22.4403 | -106.1326
22.444 | -106.1372
22.4477 | -106.1432
22.4514 | -106.1479
22.4552 | -106.1498
22.4589 | -106.1529
22.4626 | -106.1579
22.4663 | -106.1639
22.4701 | -106.1666
22.4738 | -106.1685
22.4775 | -106.1728
22.4812 | -106.1764
22.485 | -106.1774
22.4887 | -106.1777
22.4924 | -106.1744
22.4962 | -106.1732
22.4999 | -106.1699
22.5036 | -106.1602
22.5073 | -106.1388
22.5111 | -106.108
22.5148 | -106.0717
22.5185 | -106.0335
22.5222 | -105.9922
22.526 | -105.9521
22.5297 | -105.9129
22.5334 | -105.8617
22.5372 | -105.8061
22.5409 | -105.7574
22.5446 | -105.7046
22.5483 | -105.6472
22.5521 | -105.5881
22.5558 | -105.5269
22.5595 | -105.4604
22.5632 | -105.3901
22.567 | -105.316
22.5707 | -105.2393
22.5744 | -105.16
22.5782 | -105.079
22.5819 | -104.968
22.5856 | -104.8834
22.5893 | -104.7989
22.5931 | -104.7127
22.5968 | -104.6268
22.6005 | -104.5402
22.6042 | -104.4529
22.608 | -104.3663
22.6117 | -104.2822
22.6154 | -104.2
22.6191 | -104.1205
22.6229 | -104.0449
22.6266 | -103.9719
22.6303 | -103.9019
22.6341 | -103.8349
22.6378 | -103.771
22.6415 | -103.7123
22.6452 | -103.6585
22.649 | -103.6102
22.6527 | -103.5666
22.6564 | -103.5276
22.6601 | -103.4945
22.6639 | -103.4665
22.6676 | -103.4431
22.6713 | -103.4266
22.6751 | -103.4157
22.6788 | -103.4109
22.6825 | -103.4119
22.6862 | -103.4189
22.69 | -103.4306
22.6937 | -103.4473
22.6974 | -103.4698
22.7011 | -103.4992
22.7049 | -103.5324
22.7086 | -103.5708
22.7123 | -103.6141
22.716 | -103.6625
22.7198 | -103.7146
22.7235 | -103.7707
22.7272 | -103.8302
22.731 | -103.8921
22.7347 | -103.957
22.7384 | -104.0235
22.7421 | -104.0923
22.7459 | -104.1646
22.7496 | -104.2379
22.7533 | -104.3112
22.757 | -104.3865
22.7608 | -104.4606
22.7645 | -104.5333
22.7682 | -104.6065
22.772 | -104.6791
22.7757 | -104.7504
22.7794 | -104.8195
22.7831 | -104.8862
22.7869 | -104.9505
22.7906 | -105.0142
22.7943 | -105.076
22.798 | -105.131
22.8018 | -105.1839
22.8055 | -105.237
22.8092 | -105.2855
22.813 | -105.3338
22.8167 | -105.3769
22.8204 | -105.4176
22.8241 | -105.459
22.8279 | -105.5034
22.8316 | -105.5478
22.8353 | -105.5854
22.839 | -105.6159
22.8428 | -105.643
22.8465 | -105.6603
22.8502 | -105.6703
22.8539 | -105.685
22.8577 | -105.6966
22.8614 | -105.6948
22.8651 | -105.6893
22.8689 | -105.6849
22.8726 | -105.6785
22.8763 | -105.6699
22.88 | -105.6614
22.8838 | -105.653
22.8875 | -105.6432
22.8912 | -105.6328
22.8949 | -105.6235
22.8987 | -105.614
22.9024 | -105.6044
22.9061 | -105.5961
22.9099 | -105.5884
22.9136 | -105.5791
22.9173 | -105.5692
22.921 | -105.558
22.9248 | -105.5483
22.9285 | -105.5402
22.9322 | -105.5308
22.9359 | -105.5213
22.9397 | -105.5119
22.9434 | -105.503
22.9471 | -105.4945
22.9508 | -105.4861
22.9546 | -105.4773
22.9583 | -105.4708
22.962 | -105.4649
22.9658 | -105.4583
22.9695 | -105.451
22.9732 | -105.4445
22.9769 | -105.4394
22.9807 | -105.4352
22.9844 | -105.4289
22.9881 | -105.424
22.9918 | -105.4229
22.9956 | -105.4204
22.9993 | -105.4162
23.003 | -105.4127
23.0068 | -105.4104
23.0105 | -105.4092
23.0142 | -105.4091
23.0179 | -105.4093
23.0217 | -105.4097
23.0254 | -105.4095
23.0291 | -105.4102
23.0328 | -105.4115
23.0366 | -105.412
23.0403 | -105.4136
23.044 | -105.4173
23.0478 | -105.42
23.0515 | -105.423
23.0552 | -105.4281
23.0589 | -105.4323
23.0627 | -105.4357
23.0664 | -105.4399
23.0701 | -105.4435
23.0738 | -105.4464
23.0776 | -105.449
23.0813 | -105.4502
23.085 | -105.4536
23.0887 | -105.4591
23.0925 | -105.4603
23.0962 | -105.4609
23.0999 | -105.4624
23.1037 | -105.4645
23.1074 | -105.4666
23.1111 | -105.4682
23.1148 | -105.469
23.1186 | -105.4682
23.1223 | -105.4689
23.126 | -105.4715
23.1297 | -105.4734
23.1335 | -105.4732
23.1372 | -105.4724
23.1409 | -105.4713
23.1447 | -105.4701
23.1484 | -105.4697
23.1521 | -105.4658
23.1558 | -105.4628
23.1596 | -105.4604
23.1633 | -105.4595
23.167 | -105.4583
23.1707 | -105.4551
23.1745 | -105.4521
23.1782 | -105.4483
23.1819 | -105.4445
23.1857 | -105.4399
23.1894 | -105.4377
23.1931 | -105.4377
23.1968 | -105.4365
23.2006 | -105.4334
23.2043 | -105.4304
23.208 | -105.4287
23.2117 | -105.427
23.2155 | -105.4267
23.2192 | -105.4265
23.2229 | -105.4264
23.2266 | -105.4251
23.2304 | -105.4223
23.2341 | -105.4216
23.2378 | -105.423
23.2416 | -105.4244
23.2453 | -105.4265
23.249 | -105.4296
23.2527 | -105.4317
23.2565 | -105.4338
23.2602 | -105.4361
23.2639 | -105.438
23.2676 | -105.4412
23.2714 | -105.445
23.2751 | -105.4479
23.2788 | -105.4509
23.2826 | -105.4563
23.2863 | -105.4599
23.29 | -105.4627
23.2937 | -105.4685
23.2975 | -105.4748
23.3012 | -105.4779
23.3049 | -105.4814
23.3086 | -105.4849
23.3124 | -105.4915
23.3161 | -105.4985
23.3198 | -105.504
23.3236 | -105.5096
23.3273 | -105.515
23.331 | -105.5193
23.3347 | -105.5228
23.3385 | -105.5278
23.3422 | -105.5347
23.3459 | -105.5401
23.3496 | -105.5455
23.3534 | -105.5495
23.3571 | -105.5525
23.3608 | -105.5559
23.3645 | -105.5614
23.3683 | -105.5652
23.372 | -105.5692
23.3757 | -105.5752
23.3795 | -105.5818
23.3832 | -105.5872
23.3869 | -105.594
23.3906 | -105.6014
23.3944 | -105.6077
23.3981 | -105.6148
23.4018 | -105.6244
23.4055 | -105.6331
23.4093 | -105.6422
23.413 | -105.65
23.4167 | -105.6575
23.4205 | -105.6649
23.4242 | -105.6705
23.4279 | -105.675
23.4316 | -105.6814
23.4354 | -105.6879
23.4391 | -105.6961
23.4428 | -105.7046
23.4465 | -105.712
23.4503 | -105.7206
23.454 | -105.7291
23.4577 | -105.7372
23.4614 | -105.7447
23.4652 | -105.7519
23.4689 | -105.7588
23.4726 | -105.7668
23.4764 | -105.7757
23.4801 | -105.7826
23.4838 | -105.7907
23.4875 | -105.7964
23.4913 | -105.8022
23.495 | -105.809
23.4987 | -105.8166
23.5024 | -105.8243
23.5062 | -105.8295
23.5099 | -105.8329
23.5136 | -105.8361
23.5174 | -105.8399
23.5211 | -105.8417
23.5248 | -105.8427
23.5285 | -105.8419
23.5323 | -105.8415
23.536 | -105.8438
23.5397 | -105.845
23.5434 | -105.8454
23.5472 | -105.8447
23.5509 | -105.8456
23.5546 | -105.8442
23.5584 | -105.8438
23.5621 | -105.8444
23.5658 | -105.8424
23.5695 | -105.8392
23.5733 | -105.8374
23.577 | -105.8364
23.5807 | -105.8361
23.5844 | -105.8357
23.5882 | -105.8348
23.5919 | -105.8345
23.5956 | -105.8346
23.5993 | -105.834
23.6031 | -105.8333
23.6068 | -105.8332
23.6105 | -105.8324
23.6143 | -105.8314
23.618 | -105.8318
23.6217 | -105.8346
23.6254 | -105.836
23.6292 | -105.8386
23.6329 | -105.8425
23.6366 | -105.845
23.6403 | -105.8485
23.6441 | -105.8512
23.6478 | -105.8548
23.6515 | -105.8607
23.6553 | -105.8722
23.659 | -105.8801
23.6627 | -105.889
23.6664 | -105.8991
23.6702 | -105.9104
23.6739 | -105.9204
23.6776 | -105.9301
23.6813 | -105.9406
23.6851 | -105.951
23.6888 | -105.964
23.6925 | -105.9785
23.6962 | -105.9923
23.7 | -106.0069
23.7037 | -106.0221
23.7074 | -106.0339
23.7112 | -106.0458
23.7149 | -106.059
23.7186 | -106.0729
23.7223 | -106.0858
23.7261 | -106.0988
23.7298 | -106.1113
23.7335 | -106.1249
23.7372 | -106.1384
23.741 | -106.1527
23.7447 | -106.1656
23.7484 | -106.18
23.7522 | -106.1941
23.7559 | -106.2098
23.7596 | -106.2271
23.7633 | -106.2441
23.7671 | -106.2572
23.7708 | -106.2689
23.7745 | -106.2802
23.7782 | -106.2912
23.782 | -106.3025
23.7857 | -106.312
23.7894 | -106.3216
23.7932 | -106.3299
23.7969 | -106.3394
23.8006 | -106.3485
23.8043 | -106.3582
23.8081 | -106.3656
23.8118 | -106.3734
23.8155 | -106.3814
23.8192 | -106.3895
23.823 | -106.3973
23.8267 | -106.4031
23.8304 | -106.4054
23.8341 | -106.4062
23.8379 | -106.4082
23.8416 | -106.4129
23.8453 | -106.418
23.8491 | -106.4218
23.8528 | -106.4237
23.8565 | -106.4256
23.8602 | -106.4253
23.864 | -106.4251
23.8677 | -106.4268
23.8714 | -106.4247
23.8751 | -106.4232
23.8789 | -106.42
23.8826 | -106.418
23.8863 | -106.4147
23.8901 | -106.4142
23.8938 | -106.4137
23.8975 | -106.4113
23.9012 | -106.4093
23.905 | -106.4075
23.9087 | -106.4053
23.9124 | -106.4012
23.9161 | -106.3973
23.9199 | -106.3952
23.9236 | -106.3926
23.9273 | -106.3904
23.9311 | -106.3866
23.9348 | -106.3827
23.9385 | -106.3803
23.9422 | -106.3787
23.946 | -106.3767
23.9497 | -106.3755
23.9534 | -106.3747
23.9571 | -106.3702
23.9609 | -106.3664
23.9646 | -106.3636
23.9683 | -106.3612
23.972 | -106.3604
23.9758 | -106.3595
23.9795 | -106.3597
23.9832 | -106.3588
23.987 | -106.3583
23.9907 | -106.358
23.9944 | -106.357
23.9981 | -106.3537
24.0019 | -106.3526
24.0056 | -106.3527
24.0093 | -106.3517
24.013 | -106.3502
24.0168 | -106.3507
24.0205 | -106.3522
24.0242 | -106.353
24.028 | -106.3532
24.0317 | -106.3536
24.0354 | -106.3541
24.0391 | -106.3537
24.0429 | -106.3543
24.0466 | -106.3564
24.0503 | -106.3599
24.054 | -106.3624
24.0578 | -106.3643
24.0615 | -106.3644
24.0652 | -106.3651
24.0689 | -106.3661
24.0727 | -106.3658
24.0764 | -106.3648
24.0801 | -106.364
24.0839 | -106.3658
24.0876 | -106.3658
24.0913 | -106.3657
24.095 | -106.3664
24.0988 | -106.3671
24.1025 | -106.3666
24.1062 | -106.3667
24.1099 | -106.3689
24.1137 | -106.3692
24.1174 | -106.3683
24.1211 | -106.3666
24.1249 | -106.3657
24.1286 | -106.3649
24.1323 | -106.3649
24.136 | -106.3638
24.1398 | -106.3635
24.1435 | -106.3603
24.1472 | -106.3569
24.1509 | -106.354
24.1547 | -106.3522
24.1584 | -106.3491
24.1621 | -106.3442
24.1659 | -106.3389
24.1696 | -106.3332
24.1733 | -106.3266
24.177 | -106.3187
24.1808 | -106.3102
24.1845 | -106.3038
24.1882 | -106.2969
24.1919 | -106.2899
24.1957 | -106.2805
24.1994 | -106.2732
24.2031 | -106.2664
24.2068 | -106.2589
24.2106 | -106.2507
24.2143 | -106.2439
24.218 | -106.2347
24.2218 | -106.2262
24.2255 | -106.2183
24.2292 | -106.211
24.2329 | -106.2011
24.2367 | -106.1906
24.2404 | -106.1818
24.2441 | -106.1735
24.2478 | -106.1667
24.2516 | -106.1595
24.2553 | -106.153
24.259 | -106.1458
24.2628 | -106.1382
24.2665 | -106.1318
24.2702 | -106.1255
24.2739 | -106.1192
24.2777 | -106.1089
24.2814 | -106.0945
24.2851 | -106.0696
24.2888 | -106.0469
24.2926 | -106.0158
24.2963 | -105.9756
24.3 | -105.9394
24.3038 | -105.9096
24.3075 | -105.8847
24.3112 | -105.861
24.3149 | -105.8359
24.3187 | -105.8095
24.3224 | -105.7813
24.3261 | -105.7521
24.3298 | -105.721
24.3336 | -105.6891
24.3373 | -105.6543
24.341 | -105.6196
24.3447 | -105.583
24.3485 | -105.5435
24.3522 | -105.5031
24.3559 | -105.4624
24.3597 | -105.4192
24.3634 | -105.3753
24.3671 | -105.33
24.3708 | -105.2848
24.3746 | -105.2386
24.3783 | -105.1908
24.382 | -105.1438
24.3857 | -105.0974
24.3895 | -105.0507
24.3932 | -105.0032
24.3969 | -104.957
24.4007 | -104.9118
24.4044 | -104.8675
24.4081 | -104.8242
24.4118 | -104.7827
24.4156 | -104.7423
24.4193 | -104.7042
24.423 | -104.6689
24.4267 | -104.6347
24.4305 | -104.6032
24.4342 | -104.5743
24.4379 | -104.5483
24.4416 | -104.5255
24.4454 | -104.5064
24.4491 | -104.4914
24.4528 | -104.4779
24.4566 | -104.4679
24.4603 | -104.4632
24.464 | -104.463
24.4677 | -104.4655
24.4715 | -104.4717
24.4752 | -104.4816
24.4789 | -104.4941
24.4826 | -104.5079
24.4864 | -104.5267
24.4901 | -104.548
24.4938 | -104.5691
24.4976 | -104.5943
24.5013 | -104.6241
24.505 | -104.6561
24.5087 | -104.6911
24.5125 | -104.7286
24.5162 | -104.7682
24.5199 | -104.8086
24.5236 | -104.8495
24.5274 | -104.8908
24.5311 | -104.9328
24.5348 | -104.9758
24.5386 | -105.0203
24.5423 | -105.064
24.546 | -105.1068
24.5497 | -105.149
24.5535 | -105.1916
24.5572 | -105.2307
24.5609 | -105.2662
24.5646 | -105.2978
24.5684 | -105.3217
24.5721 | -105.334
24.5758 | -105.3287
24.5795 | -105.3119
24.5833 | -105.2863
24.587 | -105.2555
24.5907 | -105.2269
24.5945 | -105.2025
24.5982 | -105.174
24.6019 | -105.1459
24.6056 | -105.116
24.6094 | -105.0775
24.6131 | -105.0334
24.6168 | -104.9847
24.6205 | -104.9305
24.6243 | -104.8747
24.628 | -104.8167
24.6317 | -104.76
24.6355 | -104.6959
24.6392 | -104.6323
24.6429 | -104.568
24.6466 | -104.497
24.6504 | -104.4152
24.6541 | -104.3256
24.6578 | -104.2332
24.6615 | -104.1405
24.6653 | -104.049
24.669 | -103.9575
24.6727 | -103.8675
24.6764 | -103.7789
24.6802 | -103.6921
24.6839 | -103.6083
24.6876 | -103.5263
24.6914 | -103.4461
24.6951 | -103.3705
24.6988 | -103.2971
24.7025 | -103.2268
24.7063 | -103.16
24.71 | -103.0979
24.7137 | -103.0407
24.7174 | -102.9864
24.7212 | -102.9379
24.7249 | -102.8949
24.7286 | -102.8557
24.7324 | -102.8629
24.7361 | -102.8356
24.7398 | -102.8137
24.7435 | -102.7965
24.7473 | -102.7847
24.751 | -102.7784
24.7547 | -102.778
24.7584 | -102.7826
24.7622 | -102.792
24.7659 | -102.8077
24.7696 | -102.8282
24.7734 | -102.8538
24.7771 | -102.8844
24.7808 | -102.9196
24.7845 | -102.9603
24.7883 | -103.0056
24.792 | -103.0548
24.7957 | -103.1081
24.7994 | -103.1645
24.8032 | -103.2263
24.8069 | -103.2919
24.8106 | -103.3602
24.8143 | -103.4308
24.8181 | -103.5047
24.8218 | -103.5812
24.8255 | -103.6593
24.8293 | -103.7398
24.833 | -103.8237
24.8367 | -103.9063
24.8404 | -103.9907
24.8442 | -104.078
24.8479 | -104.1639
24.8516 | -104.2479
24.8553 | -104.3328
24.8591 | -104.4163
24.8628 | -104.4979
24.8665 | -104.5774
24.8703 | -104.6555
24.874 | -104.7323
24.8777 | -104.8073
24.8814 | -104.879
24.8852 | -104.9474
24.8889 | -105.0135
24.8926 | -105.0773
24.8963 | -105.1383
24.9001 | -105.1957
24.9038 | -105.2501
24.9075 | -105.3034
24.9113 | -105.3537
24.915 | -105.4037
24.9187 | -105.4552
24.9224 | -105.5091
24.9262 | -105.5691
24.9299 | -105.63
24.9336 | -105.6883
24.9373 | -105.738
24.9411 | -105.7706
24.9448 | -105.7857
24.9485 | -105.7973
24.9522 | -105.8012
24.956 | -105.8022
24.9597 | -105.8072
24.9634 | -105.8116
24.9672 | -105.8146
24.9709 | -105.818
24.9746 | -105.8195
24.9783 | -105.8183
24.9821 | -105.8175
24.9858 | -105.8168
24.9895 | -105.8143
24.9932 | -105.8109
24.997 | -105.8073
25.0007 | -105.8038
25.0044 | -105.7996
25.0082 | -105.7973
25.0119 | -105.7939
25.0156 | -105.7884
25.0193 | -105.784
25.0231 | -105.7801
25.0268 | -105.7759
25.0305 | -105.7719
25.0342 | -105.7674
25.038 | -105.7618
25.0417 | -105.7563
25.0454 | -105.7519
25.0492 | -105.749
25.0529 | -105.7458
25.0566 | -105.7426
25.0603 | -105.7382
25.0641 | -105.7349
25.0678 | -105.733
25.0715 | -105.7289
25.0752 | -105.7247
25.079 | -105.7227
25.0827 | -105.7189
25.0864 | -105.7143
25.0901 | -105.7113
25.0939 | -105.709
25.0976 | -105.7087
25.1013 | -105.7068
25.1051 | -105.7064
25.1088 | -105.7059
25.1125 | -105.705
25.1162 | -105.7049
25.12 | -105.7052
25.1237 | -105.706
25.1274 | -105.7065
25.1311 | -105.7078
25.1349 | -105.7104
25.1386 | -105.7137
25.1423 | -105.7164
25.1461 | -105.719
25.1498 | -105.7233
25.1535 | -105.729
25.1572 | -105.733
25.161 | -105.7386
25.1647 | -105.7439
25.1684 | -105.7504
25.1721 | -105.7567
25.1759 | -105.7643
25.1796 | -105.7739
25.1833 | -105.7827
25.187 | -105.791
25.1908 | -105.803
25.1945 | -105.815
25.1982 | -105.8239
25.202 | -105.8321
25.2057 | -105.8417
25.2094 | -105.8524
25.2131 | -105.8628
25.2169 | -105.8729
25.2206 | -105.8848
25.2243 | -105.896
25.228 | -105.9038
25.2318 | -105.91
25.2355 | -105.9188
25.2392 | -105.9291
25.243 | -105.937
25.2467 | -105.9447
25.2504 | -105.9531
25.2541 | -105.9627
25.2579 | -105.9703
25.2616 | -105.9766
25.2653 | -105.9842
25.269 | -105.986
25.2728 | -105.9926
25.2765 | -105.9982
25.2802 | -106.0032
25.284 | -106.0073
25.2877 | -106.0085
25.2914 | -106.0095
25.2951 | -106.0102
25.2989 | -106.0109
25.3026 | -106.0109
25.3063 | -106.0123
25.31 | -106.013
25.3138 | -106.012
25.3175 | -106.0084
25.3212 | -106.0044
25.3249 | -106.0008
25.3287 | -105.9969
25.3324 | -105.9939
25.3361 | -105.9902
25.3399 | -105.9849
25.3436 | -105.9802
25.3473 | -105.9741
25.351 | -105.9693
25.3548 | -105.9635
25.3585 | -105.9586
25.3622 | -105.9541
25.3659 | -105.947
25.3697 | -105.94
25.3734 | -105.9337
25.3771 | -105.9255
25.3809 | -105.9179
25.3846 | -105.9123
25.3883 | -105.9061
25.392 | -105.8982
25.3958 | -105.8931
25.3995 | -105.8892
25.4032 | -105.8827
25.4069 | -105.8755
25.4107 | -105.8688
25.4144 | -105.8641
25.4181 | -105.8604
25.4218 | -105.8529
25.4256 | -105.8463
25.4293 | -105.842
25.433 | -105.8384
25.4368 | -105.8361
25.4405 | -105.836
25.4442 | -105.8347
25.4479 | -105.8312
25.4517 | -105.8294
25.4554 | -105.8279
25.4591 | -105.8269
25.4628 | -105.8245
25.4666 | -105.8219
25.4703 | -105.8222
25.474 | -105.8229
25.4778 | -105.8206
25.4815 | -105.8195
25.4852 | -105.8205
25.4889 | -105.821
25.4927 | -105.8227
25.4964 | -105.8225
25.5001 | -105.8219
25.5038 | -105.8234
25.5076 | -105.8253
25.5113 | -105.8265
25.515 | -105.8281
25.5188 | -105.8279
25.5225 | -105.8274
25.5262 | -105.8275
25.5299 | -105.8298
25.5337 | -105.8309
25.5374 | -105.8306
25.5411 | -105.8312
25.5448 | -105.831
25.5486 | -105.8288
25.5523 | -105.8275
25.556 | -105.8275
25.5597 | -105.8272
25.5635 | -105.8284
25.5672 | -105.8281
25.5709 | -105.8283
25.5747 | -105.8299
25.5784 | -105.8288
25.5821 | -105.8276
25.5858 | -105.8272
25.5896 | -105.8262
25.5933 | -105.8245
25.597 | -105.8209
25.6007 | -105.8189
25.6045 | -105.8174
25.6082 | -105.8148
25.6119 | -105.8119
25.6157 | -105.8101
25.6194 | -105.8074
25.6231 | -105.8044
25.6268 | -105.8007
25.6306 | -105.7975
25.6343 | -105.7923
25.638 | -105.7862
25.6417 | -105.7808
25.6455 | -105.7766
25.6492 | -105.7722
25.6529 | -105.765
25.6566 | -105.7581
25.6604 | -105.7538
25.6641 | -105.7506
25.6678 | -105.7455
25.6716 | -105.7411
25.6753 | -105.7368
25.679 | -105.7341
25.6827 | -105.731
25.6865 | -105.7278
25.6902 | -105.724
25.6939 | -105.7186
25.6976 | -105.7147
25.7014 | -105.7114
25.7051 | -105.7082
25.7088 | -105.7052
25.7126 | -105.7003
25.7163 | -105.6956
25.72 | -105.6919
25.7237 | -105.6891
25.7275 | -105.6856
25.7312 | -105.6825
25.7349 | -105.6796
25.7386 | -105.6779
25.7424 | -105.6764
25.7461 | -105.6745
25.7498 | -105.6737
25.7536 | -105.6733
25.7573 | -105.6729
25.761 | -105.6718
25.7647 | -105.6706
25.7685 | -105.671
25.7722 | -105.6727
25.7759 | -105.6743
25.7796 | -105.6751
25.7834 | -105.6801
25.7871 | -105.6841
25.7908 | -105.6853
25.7945 | -105.6862
25.7983 | -105.6863
25.802 | -105.6857
25.8057 | -105.6928
25.8095 | -105.6971
25.8132 | -105.7006
25.8169 | -105.7024
25.8206 | -105.7051
25.8244 | -105.7084
25.8281 | -105.7126
25.8318 | -105.7179
25.8355 | -105.722
25.8393 | -105.7251
25.843 | -105.729
25.8467 | -105.7345
25.8505 | -105.7384
25.8542 | -105.7411
25.8579 | -105.7449
25.8616 | -105.7508
25.8654 | -105.7576
25.8691 | -105.7627
25.8728 | -105.7656
25.8765 | -105.769
25.8803 | -105.7742
25.884 | -105.7803
25.8877 | -105.787
25.8915 | -105.7931
25.8952 | -105.7961
25.8989 | -105.7975
25.9026 | -105.8
25.9064 | -105.8028
25.9101 | -105.8064
25.9138 | -105.8103
25.9175 | -105.8129
25.9213 | -105.8148
25.925 | -105.8153
25.9287 | -105.8139
25.9324 | -105.8138
25.9362 | -105.8113
25.9399 | -105.8077
25.9436 | -105.8073
25.9474 | -105.8071
25.9511 | -105.8062
25.9548 | -105.8063
25.9585 | -105.8033
25.9623 | -105.8012
25.966 | -105.7972
25.9697 | -105.7936
25.9734 | -105.7911
25.9772 | -105.7876
25.9809 | -105.7842
25.9846 | -105.7822
25.9884 | -105.7805
25.9921 | -105.7779
25.9958 | -105.7748
25.9995 | -105.7739
26.0033 | -105.7744
26.007 | -105.7757
26.0107 | -105.7762
26.0144 | -105.7775
26.0182 | -105.7808
26.0219 | -105.7818
26.0256 | -105.7807
26.0294 | -105.7802
26.0331 | -105.7798
26.0368 | -105.7783
26.0405 | -105.7778
26.0443 | -105.7788
26.048 | -105.7805
26.0517 | -105.7818
26.0554 | -105.7833
26.0592 | -105.7866
26.0629 | -105.7903
26.0666 | -105.7933
26.0703 | -105.7943
26.0741 | -105.7965
26.0778 | -105.8007
26.0815 | -105.8049
26.0853 | -105.808
26.089 | -105.8117
26.0927 | -105.8154
26.0964 | -105.8192
26.1002 | -105.8227
26.1039 | -105.8259
26.1076 | -105.8274
26.1113 | -105.8256
26.1151 | -105.8245
26.1188 | -105.8259
26.1225 | -105.8278
26.1263 | -105.829
26.13 | -105.8314
26.1337 | -105.8334
26.1374 | -105.832
26.1412 | -105.8311
26.1449 | -105.8326
26.1486 | -105.8338
26.1523 | -105.8347
26.1561 | -105.835
26.1598 | -105.8352
26.1635 | -105.8333
26.1672 | -105.8335
26.171 | -105.8292
26.1747 | -105.8247
26.1784 | -105.8211
26.1822 | -105.8179
26.1859 | -105.815
26.1896 | -105.811
26.1933 | -105.8065
26.1971 | -105.8024
26.2008 | -105.7969
26.2045 | -105.792
26.2082 | -105.7881
26.212 | -105.784
26.2157 | -105.7782
26.2194 | -105.7713
26.2232 | -105.7637
26.2269 | -105.7561
26.2306 | -105.7486
26.2343 | -105.7422
26.2381 | -105.7368
26.2418 | -105.732
26.2455 | -105.7261
26.2492 | -105.7181
26.253 | -105.71
26.2567 | -105.7032
26.2604 | -105.6972
26.2642 | -105.6922
26.2679 | -105.6884
26.2716 | -105.6839
26.2753 | -105.6801
26.2791 | -105.6757
26.2828 | -105.6719
26.2865 | -105.6687
26.2902 | -105.6675
26.294 | -105.6683
26.2977 | -105.6679
26.3014 | -105.6674
26.3051 | -105.6701
26.3089 | -105.673
26.3126 | -105.6733
26.3163 | -105.6755
26.3201 | -105.6782
26.3238 | -105.6806
26.3275 | -105.6849
26.3312 | -105.6892
26.335 | -105.6951
26.3387 | -105.6997
26.3424 | -105.7037
26.3461 | -105.7113
26.3499 | -105.7152
26.3536 | -105.7205
26.3573 | -105.7262
26.3611 | -105.7322
26.3648 | -105.7409
26.3685 | -105.7505
26.3722 | -105.76
26.376 | -105.7678
26.3797 | -105.7744
26.3834 | -105.7827
26.3871 | -105.7925
26.3909 | -105.8006
26.3946 | -105.8095
26.3983 | -105.8201
26.402 | -105.8295
26.4058 | -105.8383
26.4095 | -105.8485
26.4132 | -105.8586
26.417 | -105.8669
26.4207 | -105.8757
26.4244 | -105.8851
26.4281 | -105.894
26.4319 | -105.9019
26.4356 | -105.9105
26.4393 | -105.9194
26.443 | -105.9272
26.4468 | -105.9344
26.4505 | -105.9416
26.4542 | -105.9492
26.458 | -105.957
26.4617 | -105.962
26.4654 | -105.969
26.4691 | -105.9755
26.4729 | -105.9805
26.4766 | -105.9855
26.4803 | -105.9913
26.484 | -105.998
26.4878 | -106.0046
26.4915 | -106.009
26.4952 | -106.0109
26.499 | -106.015
26.5027 | -106.0177
26.5064 | -106.0203
26.5101 | -106.0259
26.5139 | -106.0292
26.5176 | -106.0286
26.5213 | -106.0285
26.525 | -106.0302
26.5288 | -106.0329
26.5325 | -106.0348
26.5362 | -106.0373
26.5399 | -106.0398
26.5437 | -106.0434
26.5474 | -106.0464
26.5511 | -106.0494
26.5549 | -106.0521
26.5586 | -106.055
26.5623 | -106.0574
26.566 | -106.0588
26.5698 | -106.0604
26.5735 | -106.0628
26.5772 | -106.0636
26.5809 | -106.0639
26.5847 | -106.064
26.5884 | -106.0661
26.5921 | -106.069
26.5959 | -106.0715
26.5996 | -106.0728
26.6033 | -106.0734
26.607 | -106.0744
26.6108 | -106.0746
26.6145 | -106.0754
26.6182 | -106.0767
26.6219 | -106.0789
26.6257 | -106.0802
26.6294 | -106.0792
26.6331 | -106.077
26.6369 | -106.0756
26.6406 | -106.0754
26.6443 | -106.0754
26.648 | -106.0742
26.6518 | -106.0743
26.6555 | -106.0744
26.6592 | -106.0726
26.6629 | -106.0708
26.6667 | -106.0687
26.6704 | -106.0663
26.6741 | -106.0639
26.6778 | -106.0613
26.6816 | -106.0592
26.6853 | -106.0582
26.689 | -106.0557
26.6928 | -106.0535
26.6965 | -106.0513
26.7002 | -106.0484
26.7039 | -106.0447
26.7077 | -106.0404
26.7114 | -106.0386
26.7151 | -106.038
26.7188 | -106.0362
26.7226 | -106.0342
26.7263 | -106.0324
26.73 | -106.0293
26.7338 | -106.0259
26.7375 | -106.0218
26.7412 | -106.0177
26.7449 | -106.0137
26.7487 | -106.0077
26.7524 | -106.0046
26.7561 | -106.0033
26.7598 | -106.0007
26.7636 | -105.9985
26.7673 | -105.9985
26.771 | -105.9986
26.7748 | -105.9969
26.7785 | -105.9954
26.7822 | -105.9945
26.7859 | -105.9911
26.7897 | -105.9883
26.7934 | -105.9872
26.7971 | -105.987
26.8008 | -105.986
26.8046 | -105.9862
26.8083 | -105.9871
26.812 | -105.9877
26.8157 | -105.9888
26.8195 | -105.9907
26.8232 | -105.992
26.8269 | -105.9924
26.8307 | -105.9928
26.8344 | -105.9931
26.8381 | -105.9946
26.8418 | -105.9952
26.8456 | -105.9952
26.8493 | -105.9966
26.853 | -105.9983
26.8567 | -106.0004
26.8605 | -106.0039
26.8642 | -106.0052
26.8679 | -106.0064
26.8717 | -106.0099
26.8754 | -106.0125
26.8791 | -106.0149
26.8828 | -106.0181
26.8866 | -106.0213
26.8903 | -106.0237
26.894 | -106.0274
26.8977 | -106.0321
26.9015 | -106.0362
26.9052 | -106.0399
26.9089 | -106.0443
26.9126 | -106.0507
26.9164 | -106.0561
26.9201 | -106.0594
26.9238 | -106.0626
26.9276 | -106.0658
26.9313 | -106.0684
26.935 | -106.0723
26.9387 | -106.076
26.9425 | -106.0805
26.9462 | -106.0863
26.9499 | -106.0911
26.9536 | -106.0957
26.9574 | -106.1002
26.9611 | -106.105
26.9648 | -106.1123
26.9686 | -106.1168
26.9723 | -106.1212
26.976 | -106.1273
26.9797 | -106.1335
26.9835 | -106.1388
26.9872 | -106.1423
26.9909 | -106.146
26.9946 | -106.1496
26.9984 | -106.1532
27.0021 | -106.1551
27.0058 | -106.1597
27.0096 | -106.1632
27.0133 | -106.1669
27.017 | -106.1711
27.0207 | -106.1741
27.0245 | -106.1753
27.0282 | -106.1752
27.0319 | -106.1761
27.0356 | -106.1768
27.0394 | -106.1776
27.0431 | -106.1789
27.0468 | -106.1805
27.0505 | -106.1829
27.0543 | -106.1844
27.058 | -106.1849
27.0617 | -106.1857
27.0655 | -106.1851
27.0692 | -106.1839
27.0729 | -106.1839
27.0766 | -106.1828
27.0804 | -106.1823
27.0841 | -106.1839
27.0878 | -106.1847
27.0915 | -106.1831
27.0953 | -106.182
27.099 | -106.1802
27.1027 | -106.18
27.1065 | -106.1812
27.1102 | -106.1827
27.1139 | -106.1831
27.1176 | -106.1817
27.1214 | -106.179
27.1251 | -106.1788
27.1288 | -106.1805
27.1325 | -106.1807
27.1363 | -106.1786
27.14 | -106.1756
27.1437 | -106.1734
27.1474 | -106.1702
27.1512 | -106.1687
27.1549 | -106.1685
27.1586 | -106.1678
27.1624 | -106.1671
27.1661 | -106.1705
27.1698 | -106.1717
27.1735 | -106.1716
27.1773 | -106.1703
27.181 | -106.1696
27.1847 | -106.169
27.1884 | -106.168
27.1922 | -106.1673
27.1959 | -106.166
27.1996 | -106.1664
27.2034 | -106.1685
27.2071 | -106.1709
27.2108 | -106.1728
27.2145 | -106.1749
27.2183 | -106.1767
27.222 | -106.1777
27.2257 | -106.1772
27.2294 | -106.1766
27.2332 | -106.1755
27.2369 | -106.1754
27.2406 | -106.1763
27.2444 | -106.1789
27.2481 | -106.1822
27.2518 | -106.1843
27.2555 | -106.1852
27.2593 | -106.1851
27.263 | -106.1864
27.2667 | -106.1875
27.2704 | -106.188
27.2742 | -106.1873
27.2779 | -106.1878
27.2816 | -106.1895
27.2853 | -106.1907
27.2891 | -106.1906
27.2928 | -106.1892
27.2965 | -106.1884
27.3003 | -106.189
27.304 | -106.1894
27.3077 | -106.1895
27.3114 | -106.1885
27.3152 | -106.1883
27.3189 | -106.1883
27.3226 | -106.1859
27.3263 | -106.1847
27.3301 | -106.1824
27.3338 | -106.1807
27.3375 | -106.1791
27.3413 | -106.1763
27.345 | -106.1724
27.3487 | -106.1688
27.3524 | -106.167
27.3562 | -106.1653
27.3599 | -106.1633
27.3636 | -106.1589
27.3673 | -106.1559
27.3711 | -106.1544
27.3748 | -106.1532
27.3785 | -106.1513
27.3822 | -106.1502
27.386 | -106.1497
27.3897 | -106.1474
27.3934 | -106.1453
27.3972 | -106.1431
27.4009 | -106.1398
27.4046 | -106.1376
27.4083 | -106.1378
27.4121 | -106.1372
27.4158 | -106.1334
27.4195 | -106.1326
27.4232 | -106.1306
27.427 | -106.1293
27.4307 | -106.1272
27.4344 | -106.126
27.4382 | -106.1258
27.4419 | -106.1253
27.4456 | -106.1251
27.4493 | -106.1241
27.4531 | -106.1245
27.4568 | -106.1262
27.4605 | -106.1258
27.4642 | -106.1271
27.468 | -106.128
27.4717 | -106.1272
27.4754 | -106.1277
27.4792 | -106.1289
27.4829 | -106.13
27.4866 | -106.132
27.4903 | -106.1325
27.4941 | -106.1322
27.4978 | -106.1309
27.5015 | -106.1316
27.5052 | -106.1335
27.509 | -106.1347
27.5127 | -106.1342
27.5164 | -106.134
27.5201 | -106.1347
27.5239 | -106.1366
27.5276 | -106.1403
27.5313 | -106.1426
27.5351 | -106.1434
27.5388 | -106.1443
27.5425 | -106.1448
27.5462 | -106.1461
27.55 | -106.1473
27.5537 | -106.1464
27.5574 | -106.1462
27.5611 | -106.1469
27.5649 | -106.1482
27.5686 | -106.1506
27.5723 | -106.1529
27.5761 | -106.1544
27.5798 | -106.1545
27.5835 | -106.154
27.5872 | -106.1554
27.591 | -106.1561
27.5947 | -106.1559
27.5984 | -106.1553
27.6021 | -106.1554
27.6059 | -106.1571
27.6096 | -106.1595
27.6133 | -106.1624
27.6171 | -106.1642
27.6208 | -106.1676
27.6245 | -106.1693
27.6282 | -106.1698
27.632 | -106.169
27.6357 | -106.1674
27.6394 | -106.1672
27.6431 | -106.1669
27.6469 | -106.1681
27.6506 | -106.167
27.6543 | -106.1662
27.658 | -106.1665
27.6618 | -106.1675
27.6655 | -106.1655
27.6692 | -106.1632
27.673 | -106.1618
27.6767 | -106.1595
27.6804 | -106.1597
27.6841 | -106.1577
27.6879 | -106.1555
27.6916 | -106.1533
27.6953 | -106.1513
27.699 | -106.1503
27.7028 | -106.1473
27.7065 | -106.1435
27.7102 | -106.1426
27.714 | -106.1401
27.7177 | -106.1379
27.7214 | -106.1351
27.7251 | -106.1309
27.7289 | -106.1281
27.7326 | -106.1267
27.7363 | -106.1228
27.74 | -106.1165
27.7438 | -106.1124
27.7475 | -106.1085
27.7512 | -106.1043
27.755 | -106.0983
27.7587 | -106.0921
27.7624 | -106.0875
27.7661 | -106.0829
27.7699 | -106.0769
27.7736 | -106.072
27.7773 | -106.0682
27.781 | -106.0647
27.7848 | -106.0597
27.7885 | -106.055
27.7922 | -106.0495
27.7959 | -106.0425
27.7997 | -106.0356
27.8034 | -106.0297
27.8071 | -106.0237
27.8109 | -106.018
27.8146 | -106.0134
27.8183 | -106.0079
27.822 | -106.0019
27.8258 | -105.996
27.8295 | -105.9898
27.8332 | -105.983
27.8369 | -105.9782
27.8407 | -105.9721
27.8444 | -105.9664
27.8481 | -105.9616
27.8519 | -105.958
27.8556 | -105.9535
27.8593 | -105.9479
27.863 | -105.9433
27.8668 | -105.9389
27.8705 | -105.9348
27.8742 | -105.93
27.8779 | -105.9264
27.8817 | -105.9227
27.8854 | -105.92
27.8891 | -105.918
27.8928 | -105.9133
27.8966 | -105.9104
27.9003 | -105.9081
27.904 | -105.9047
27.9078 | -105.9005
27.9115 | -105.8964
27.9152 | -105.8912
27.9189 | -105.8868
27.9227 | -105.8821
27.9264 | -105.8781
27.9301 | -105.8716
27.9338 | -105.8628
27.9376 | -105.855
27.9413 | -105.8497
27.945 | -105.8456
27.9488 | -105.8417
27.9525 | -105.8373
27.9562 | -105.8335
27.9599 | -105.8297
27.9637 | -105.8251
27.9674 | -105.8201
27.9711 | -105.8153
27.9748 | -105.8117
27.9786 | -105.8091
27.9823 | -105.8053
27.986 | -105.8011
27.9898 | -105.7965
27.9935 | -105.7916
27.9972 | -105.7851
28.0009 | -105.7789