## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 246 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 225
Hood | 02-92B/P | 164
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1012.4 | hPa |  | PRC PRO 01
Humidity | 53.1 | %/Hr
Internal temperature of calibration bench | 24.24 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1016 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.9 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.94 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.04 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3059 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 21/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 25/06/2018
Technical Manager | Anthony HUE | Approval |  | 28/06/2018
CEO | Robert GLEMOT | Approval |  | 28/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4313 | 57.4921
0.012209 | -35.9637 | 38.4126
0.018313 | -34.9148 | 28.4773
0.024417 | -34.4677 | 22.137
0.030521 | -34.3636 | 18.3659
0.036626 | -34.2224 | 15.2759
0.04273 | -34.1369 | 12.9582
0.048834 | -34.0816 | 11.3871
0.054938 | -34.0088 | 10.1928
0.061043 | -33.9335 | 9.0516
0.067147 | -33.9235 | 8.1753
0.073251 | -33.9253 | 7.4216
0.079355 | -33.9444 | 6.8678
0.08546 | -33.9628 | 6.4149
0.091564 | -33.9586 | 6.0479
0.097668 | -33.9231 | 5.8027
0.10377 | -33.8742 | 5.4553
0.10988 | -33.8423 | 5.0904
0.11598 | -33.8327 | 4.7264
0.12209 | -33.8497 | 4.4452
0.12819 | -33.8831 | 4.2139
0.13429 | -33.9155 | 3.9758
0.1404 | -33.9325 | 3.6993
0.1465 | -33.948 | 3.438
0.15261 | -33.9549 | 3.1947
0.15871 | -33.9511 | 2.9959
0.16482 | -33.9362 | 2.8376
0.17092 | -33.9111 | 2.78
0.17702 | -33.8916 | 2.6853
0.18313 | -33.8738 | 2.6051
0.18923 | -33.8608 | 2.5007
0.19534 | -33.8535 | 2.3982
0.20144 | -33.8523 | 2.2852
0.20754 | -33.8551 | 2.211
0.21365 | -33.8576 | 2.1358
0.21975 | -33.8621 | 2.0701
0.22586 | -33.8704 | 2.0035
0.23196 | -33.8792 | 1.9287
0.23807 | -33.8882 | 1.8491
0.24417 | -33.8903 | 1.7546
0.25027 | -33.8936 | 1.6696
0.25638 | -33.8934 | 1.5774
0.26248 | -33.8913 | 1.491
0.26859 | -33.8886 | 1.409
0.27469 | -33.8875 | 1.3242
0.2808 | -33.8931 | 1.2617
0.2869 | -33.8973 | 1.1948
0.293 | -33.9025 | 1.1391
0.29911 | -33.9086 | 1.1021
0.30521 | -33.9144 | 1.0669
0.31132 | -33.9183 | 1.0541
0.31742 | -33.9224 | 1.033
0.32353 | -33.9256 | 1.016
0.32963 | -33.9284 | 1.0019
0.33573 | -33.93 | 0.98979
0.34184 | -33.9323 | 0.9742
0.34794 | -33.9353 | 0.93942
0.35405 | -33.9384 | 0.90997
0.36015 | -33.9418 | 0.87687
0.36626 | -33.9439 | 0.84023
0.37236 | -33.9462 | 0.79051
0.37846 | -33.9482 | 0.73349
0.38457 | -33.947 | 0.67441
0.39067 | -33.9475 | 0.62424
0.39678 | -33.9457 | 0.56785
0.40288 | -33.9432 | 0.51458
0.40899 | -33.9392 | 0.46257
0.41509 | -33.9346 | 0.41014
0.42119 | -33.9295 | 0.36243
0.4273 | -33.9257 | 0.31318
0.4334 | -33.9221 | 0.26306
0.43951 | -33.919 | 0.21378
0.44561 | -33.9175 | 0.1678
0.45172 | -33.919 | 0.12982
0.45782 | -33.9187 | 0.087644
0.46392 | -33.9199 | 0.048886
0.47003 | -33.9215 | 0.013726
0.47613 | -33.9236 | -0.018326
0.48224 | -33.9257 | -0.048718
0.48834 | -33.9281 | -0.069715
0.49445 | -33.9301 | -0.094272
0.50055 | -33.9326 | -0.11627
0.50665 | -33.9343 | -0.134
0.51276 | -33.9354 | -0.15327
0.51886 | -33.9365 | -0.16998
0.52497 | -33.9355 | -0.17559
0.53107 | -33.9358 | -0.18495
0.53717 | -33.9361 | -0.1954
0.54328 | -33.9359 | -0.19962
0.54938 | -33.9358 | -0.20693
0.55549 | -33.9353 | -0.21518
0.56159 | -33.9351 | -0.22567
0.5677 | -33.9346 | -0.23547
0.5738 | -33.9341 | -0.24809
0.5799 | -33.9341 | -0.25969
0.58601 | -33.9345 | -0.27587
0.59211 | -33.9355 | -0.29839
0.59822 | -33.9361 | -0.31486
0.60432 | -33.9371 | -0.33634
0.61043 | -33.938 | -0.35751
0.61653 | -33.9391 | -0.37906
0.62263 | -33.9401 | -0.40221
0.62874 | -33.9415 | -0.42252
0.63484 | -33.9425 | -0.44369
0.64095 | -33.9437 | -0.46478
0.64705 | -33.9448 | -0.48482
0.65316 | -33.9458 | -0.50427
0.65926 | -33.9467 | -0.52678
0.66536 | -33.9461 | -0.54491
0.67147 | -33.9463 | -0.56112
0.67757 | -33.9468 | -0.57888
0.68368 | -33.9465 | -0.59871
0.68978 | -33.9466 | -0.6165
0.69589 | -33.9462 | -0.63462
0.70199 | -33.9453 | -0.6562
0.70809 | -33.9447 | -0.67616
0.7142 | -33.9442 | -0.69515
0.7203 | -33.9433 | -0.71439
0.72641 | -33.9424 | -0.73499
0.73251 | -33.9421 | -0.75548
0.73862 | -33.9421 | -0.77956
0.74472 | -33.9423 | -0.79858
0.75082 | -33.9418 | -0.8194
0.75693 | -33.9419 | -0.84269
0.76303 | -33.9418 | -0.86496
0.76914 | -33.9422 | -0.89196
0.77524 | -33.9424 | -0.91681
0.78135 | -33.9428 | -0.94007
0.78745 | -33.9433 | -0.96484
0.79355 | -33.9438 | -0.98934
0.79966 | -33.9447 | -1.0151
0.80576 | -33.9458 | -1.0456
0.81187 | -33.946 | -1.0752
0.81797 | -33.9466 | -1.106
0.82408 | -33.9472 | -1.1332
0.83018 | -33.9476 | -1.1645
0.83628 | -33.9484 | -1.1951
0.84239 | -33.9482 | -1.2232
0.84849 | -33.9481 | -1.2513
0.8546 | -33.9485 | -1.279
0.8607 | -33.9485 | -1.3084
0.86681 | -33.9485 | -1.3367
0.87291 | -33.948 | -1.3624
0.87901 | -33.9474 | -1.3846
0.88512 | -33.9472 | -1.4075
0.89122 | -33.947 | -1.4308
0.89733 | -33.9466 | -1.4498
0.90343 | -33.9463 | -1.471
0.90953 | -33.9464 | -1.4856
0.91564 | -33.9458 | -1.5013
0.92174 | -33.9454 | -1.5152
0.92785 | -33.9448 | -1.5284
0.93395 | -33.9448 | -1.5398
0.94006 | -33.9447 | -1.5531
0.94616 | -33.9445 | -1.5572
0.95226 | -33.9445 | -1.5683
0.95837 | -33.9446 | -1.5796
0.96447 | -33.945 | -1.5895
0.97058 | -33.9449 | -1.599
0.97668 | -33.9447 | -1.6065
0.98279 | -33.9447 | -1.6095
0.98889 | -33.9449 | -1.6138
0.99499 | -33.9449 | -1.6219
1.0011 | -33.9452 | -1.6262
1.0072 | -33.945 | -1.6294
1.0133 | -33.9449 | -1.6327
1.0194 | -33.9449 | -1.6375
1.0255 | -33.9445 | -1.6439
1.0316 | -33.9439 | -1.6474
1.0377 | -33.9433 | -1.6527
1.0438 | -33.9425 | -1.6566
1.0499 | -33.9409 | -1.6624
1.056 | -33.9397 | -1.6667
1.0621 | -33.9385 | -1.672
1.0682 | -33.938 | -1.6763
1.0743 | -33.9368 | -1.6836
1.0805 | -33.9361 | -1.6898
1.0866 | -33.9344 | -1.7002
1.0927 | -33.933 | -1.7071
1.0988 | -33.9321 | -1.7172
1.1049 | -33.9306 | -1.7243
1.111 | -33.9289 | -1.7349
1.1171 | -33.9275 | -1.7438
1.1232 | -33.9261 | -1.7577
1.1293 | -33.9249 | -1.7692
1.1354 | -33.9238 | -1.7812
1.1415 | -33.9226 | -1.7949
1.1476 | -33.9211 | -1.8085
1.1537 | -33.9197 | -1.8214
1.1598 | -33.9195 | -1.8395
1.1659 | -33.9189 | -1.8511
1.172 | -33.9181 | -1.864
1.1781 | -33.9173 | -1.8777
1.1842 | -33.9167 | -1.893
1.1903 | -33.9174 | -1.9079
1.1964 | -33.9172 | -1.9221
1.2025 | -33.9171 | -1.9371
1.2086 | -33.9176 | -1.9525
1.2147 | -33.9178 | -1.9672
1.2209 | -33.9184 | -1.9803
1.227 | -33.9203 | -1.9951
1.2331 | -33.9213 | -2.0091
1.2392 | -33.9222 | -2.0247
1.2453 | -33.9234 | -2.0364
1.2514 | -33.9247 | -2.0482
1.2575 | -33.9263 | -2.062
1.2636 | -33.9286 | -2.0732
1.2697 | -33.9304 | -2.0845
1.2758 | -33.9319 | -2.0984
1.2819 | -33.9337 | -2.1116
1.288 | -33.9354 | -2.1251
1.2941 | -33.9375 | -2.1355
1.3002 | -33.9395 | -2.1434
1.3063 | -33.9411 | -2.1531
1.3124 | -33.943 | -2.1627
1.3185 | -33.945 | -2.1731
1.3246 | -33.9468 | -2.1836
1.3307 | -33.9489 | -2.1927
1.3368 | -33.9508 | -2.1999
1.3429 | -33.9524 | -2.2083
1.349 | -33.9539 | -2.2165
1.3551 | -33.9556 | -2.2271
1.3613 | -33.9572 | -2.2346
1.3674 | -33.9585 | -2.2428
1.3735 | -33.9597 | -2.2505
1.3796 | -33.961 | -2.2575
1.3857 | -33.962 | -2.2642
1.3918 | -33.963 | -2.2736
1.3979 | -33.9643 | -2.2817
1.404 | -33.965 | -2.2915
1.4101 | -33.9659 | -2.3001
1.4162 | -33.9667 | -2.3084
1.4223 | -33.9675 | -2.3184
1.4284 | -33.9683 | -2.3264
1.4345 | -33.9689 | -2.3364
1.4406 | -33.9693 | -2.3473
1.4467 | -33.97 | -2.3564
1.4528 | -33.9706 | -2.3654
1.4589 | -33.9714 | -2.3756
1.465 | -33.9716 | -2.3831
1.4711 | -33.9725 | -2.3918
1.4772 | -33.973 | -2.4037
1.4833 | -33.9736 | -2.4158
1.4894 | -33.9744 | -2.427
1.4955 | -33.975 | -2.4383
1.5016 | -33.9757 | -2.4501
1.5078 | -33.9761 | -2.4645
1.5139 | -33.9767 | -2.4769
1.52 | -33.9772 | -2.49
1.5261 | -33.9776 | -2.5014
1.5322 | -33.9778 | -2.5138
1.5383 | -33.978 | -2.5258
1.5444 | -33.9786 | -2.5389
1.5505 | -33.979 | -2.552
1.5566 | -33.9795 | -2.5646
1.5627 | -33.9801 | -2.5778
1.5688 | -33.9808 | -2.5908
1.5749 | -33.981 | -2.6029
1.581 | -33.9814 | -2.6194
1.5871 | -33.9819 | -2.6349
1.5932 | -33.9823 | -2.6476
1.5993 | -33.9826 | -2.6642
1.6054 | -33.9829 | -2.679
1.6115 | -33.9832 | -2.6958
1.6176 | -33.9836 | -2.712
1.6237 | -33.9838 | -2.7263
1.6298 | -33.984 | -2.7415
1.6359 | -33.9844 | -2.7558
1.642 | -33.9848 | -2.7701
1.6482 | -33.9851 | -2.787
1.6543 | -33.9856 | -2.8036
1.6604 | -33.986 | -2.8181
1.6665 | -33.986 | -2.8328
1.6726 | -33.9863 | -2.8489
1.6787 | -33.9868 | -2.8657
1.6848 | -33.9866 | -2.882
1.6909 | -33.9869 | -2.8981
1.697 | -33.9871 | -2.9141
1.7031 | -33.9871 | -2.9296
1.7092 | -33.9871 | -2.9443
1.7153 | -33.9873 | -2.9604
1.7214 | -33.9868 | -2.9755
1.7275 | -33.9866 | -2.9912
1.7336 | -33.9863 | -3.0072
1.7397 | -33.9863 | -3.0223
1.7458 | -33.9864 | -3.0387
1.7519 | -33.9861 | -3.0547
1.758 | -33.986 | -3.0694
1.7641 | -33.986 | -3.0837
1.7702 | -33.9858 | -3.0992
1.7763 | -33.9858 | -3.1127
1.7824 | -33.9855 | -3.1287
1.7885 | -33.9853 | -3.1425
1.7947 | -33.9854 | -3.1572
1.8008 | -33.9852 | -3.1698
1.8069 | -33.9851 | -3.1832
1.813 | -33.9851 | -3.1972
1.8191 | -33.9847 | -3.2113
1.8252 | -33.9845 | -3.225
1.8313 | -33.9845 | -3.2392
1.8374 | -33.9844 | -3.2525
1.8435 | -33.9845 | -3.2644
1.8496 | -33.9841 | -3.2787
1.8557 | -33.984 | -3.2933
1.8618 | -33.9839 | -3.3049
1.8679 | -33.9839 | -3.3177
1.874 | -33.9839 | -3.3306
1.8801 | -33.9837 | -3.3434
1.8862 | -33.9838 | -3.3559
1.8923 | -33.9837 | -3.3679
1.8984 | -33.9833 | -3.3783
1.9045 | -33.9832 | -3.3892
1.9106 | -33.9829 | -3.4008
1.9167 | -33.9826 | -3.4118
1.9228 | -33.9825 | -3.4225
1.9289 | -33.9821 | -3.4311
1.9351 | -33.9822 | -3.4399
1.9412 | -33.9822 | -3.45
1.9473 | -33.9822 | -3.4617
1.9534 | -33.9821 | -3.4726
1.9595 | -33.9818 | -3.4839
1.9656 | -33.9813 | -3.4944
1.9717 | -33.9811 | -3.5047
1.9778 | -33.9811 | -3.516
1.9839 | -33.9811 | -3.5272
1.99 | -33.9811 | -3.5368
1.9961 | -33.9808 | -3.547
2.0022 | -33.9807 | -3.5568
2.0083 | -33.9803 | -3.5665
2.0144 | -33.9802 | -3.5762
2.0205 | -33.9802 | -3.5878
2.0266 | -33.9803 | -3.5992
2.0327 | -33.9803 | -3.6101
2.0388 | -33.9805 | -3.6238
2.0449 | -33.9806 | -3.6339
2.051 | -33.9806 | -3.6437
2.0571 | -33.9804 | -3.6549
2.0632 | -33.9803 | -3.6652
2.0693 | -33.9803 | -3.6762
2.0754 | -33.9805 | -3.6878
2.0816 | -33.9806 | -3.6996
2.0877 | -33.9805 | -3.7105
2.0938 | -33.9807 | -3.7227
2.0999 | -33.9806 | -3.7341
2.106 | -33.9808 | -3.7459
2.1121 | -33.9808 | -3.758
2.1182 | -33.9809 | -3.7678
2.1243 | -33.9811 | -3.7782
2.1304 | -33.9811 | -3.7882
2.1365 | -33.9814 | -3.7996
2.1426 | -33.9815 | -3.8128
2.1487 | -33.982 | -3.8243
2.1548 | -33.9824 | -3.8371
2.1609 | -33.9829 | -3.849
2.167 | -33.9833 | -3.8614
2.1731 | -33.9836 | -3.8731
2.1792 | -33.9843 | -3.8842
2.1853 | -33.9848 | -3.8969
2.1914 | -33.9853 | -3.9089
2.1975 | -33.9858 | -3.9209
2.2036 | -33.9862 | -3.9339
2.2097 | -33.9866 | -3.9457
2.2158 | -33.9876 | -3.9558
2.222 | -33.9881 | -3.9657
2.2281 | -33.9884 | -3.9763
2.2342 | -33.9889 | -3.9893
2.2403 | -33.9893 | -3.9991
2.2464 | -33.99 | -4.01
2.2525 | -33.9905 | -4.0206
2.2586 | -33.9913 | -4.0302
2.2647 | -33.9921 | -4.0393
2.2708 | -33.9927 | -4.0508
2.2769 | -33.9936 | -4.0621
2.283 | -33.994 | -4.0713
2.2891 | -33.9945 | -4.0814
2.2952 | -33.9951 | -4.091
2.3013 | -33.9957 | -4.101
2.3074 | -33.9964 | -4.1104
2.3135 | -33.9968 | -4.1216
2.3196 | -33.9974 | -4.1299
2.3257 | -33.9978 | -4.1402
2.3318 | -33.9981 | -4.149
2.3379 | -33.9985 | -4.1603
2.344 | -33.9989 | -4.1709
2.3501 | -33.9993 | -4.1815
2.3562 | -33.9996 | -4.1896
2.3623 | -33.9998 | -4.198
2.3685 | -34 | -4.2068
2.3746 | -34.0003 | -4.2171
2.3807 | -34.0008 | -4.2258
2.3868 | -34.0009 | -4.2346
2.3929 | -34.0013 | -4.2437
2.399 | -34.0015 | -4.2515
2.4051 | -34.002 | -4.2608
2.4112 | -34.0023 | -4.2696
2.4173 | -34.0026 | -4.2782
2.4234 | -34.0025 | -4.2866
2.4295 | -34.0028 | -4.2949
2.4356 | -34.0029 | -4.304
2.4417 | -34.0029 | -4.3139
2.4478 | -34.0029 | -4.3232
2.4539 | -34.0032 | -4.3309
2.46 | -34.0029 | -4.3404
2.4661 | -34.0029 | -4.3494
2.4722 | -34.0031 | -4.3579
2.4783 | -34.0032 | -4.3661
2.4844 | -34.0032 | -4.3744
2.4905 | -34.003 | -4.3835
2.4966 | -34.0028 | -4.3921
2.5027 | -34.0027 | -4.3999
2.5089 | -34.0026 | -4.4082
2.515 | -34.0023 | -4.4173
2.5211 | -34.0022 | -4.4254
2.5272 | -34.0019 | -4.4342
2.5333 | -34.0017 | -4.4436
2.5394 | -34.0016 | -4.4524
2.5455 | -34.0014 | -4.4616
2.5516 | -34.0013 | -4.4701
2.5577 | -34.0012 | -4.4799
2.5638 | -34.0009 | -4.4901
2.5699 | -34.0007 | -4.5001
2.576 | -34.0004 | -4.5104
2.5821 | -34.0002 | -4.5201
2.5882 | -34 | -4.5295
2.5943 | -33.9998 | -4.5391
2.6004 | -33.9996 | -4.5489
2.6065 | -33.9994 | -4.5571
2.6126 | -33.9992 | -4.5661
2.6187 | -33.9991 | -4.5746
2.6248 | -33.9989 | -4.5834
2.6309 | -33.9987 | -4.5945
2.637 | -33.9987 | -4.6056
2.6431 | -33.9987 | -4.6149
2.6492 | -33.9985 | -4.6247
2.6554 | -33.9985 | -4.6348
2.6615 | -33.9982 | -4.6445
2.6676 | -33.9982 | -4.6548
2.6737 | -33.9983 | -4.6654
2.6798 | -33.9983 | -4.6751
2.6859 | -33.9982 | -4.6855
2.692 | -33.9981 | -4.6951
2.6981 | -33.998 | -4.7061
2.7042 | -33.9981 | -4.718
2.7103 | -33.998 | -4.7279
2.7164 | -33.9982 | -4.7384
2.7225 | -33.9983 | -4.7496
2.7286 | -33.9983 | -4.7606
2.7347 | -33.9984 | -4.7704
2.7408 | -33.9986 | -4.7795
2.7469 | -33.9985 | -4.7907
2.753 | -33.9988 | -4.8012
2.7591 | -33.9989 | -4.813
2.7652 | -33.9989 | -4.8244
2.7713 | -33.9989 | -4.8341
2.7774 | -33.9993 | -4.8454
2.7835 | -33.9994 | -4.8559
2.7896 | -33.9995 | -4.865
2.7958 | -33.9999 | -4.8752
2.8019 | -34.0004 | -4.8848
2.808 | -34.0007 | -4.8965
2.8141 | -34.001 | -4.9094
2.8202 | -34.0014 | -4.9214
2.8263 | -34.0018 | -4.9322
2.8324 | -34.002 | -4.9445
2.8385 | -34.0025 | -4.9565
2.8446 | -34.003 | -4.968
2.8507 | -34.0033 | -4.98
2.8568 | -34.0038 | -4.991
2.8629 | -34.0042 | -5.0022
2.869 | -34.0046 | -5.0134
2.8751 | -34.005 | -5.0253
2.8812 | -34.0057 | -5.0384
2.8873 | -34.006 | -5.0497
2.8934 | -34.0063 | -5.0614
2.8995 | -34.0067 | -5.0733
2.9056 | -34.007 | -5.0839
2.9117 | -34.0074 | -5.0954
2.9178 | -34.0081 | -5.1079
2.9239 | -34.0085 | -5.1212
2.93 | -34.009 | -5.1345
2.9361 | -34.0095 | -5.1458
2.9423 | -34.0101 | -5.1577
2.9484 | -34.0106 | -5.1711
2.9545 | -34.0113 | -5.1843
2.9606 | -34.0118 | -5.1975
2.9667 | -34.0123 | -5.2092
2.9728 | -34.0126 | -5.2214
2.9789 | -34.0131 | -5.234
2.985 | -34.0135 | -5.2477
2.9911 | -34.014 | -5.2601
2.9972 | -34.0145 | -5.2727
3.0033 | -34.0151 | -5.2857
3.0094 | -34.0156 | -5.2986
3.0155 | -34.0162 | -5.3111
3.0216 | -34.0167 | -5.3244
3.0277 | -34.0171 | -5.3381
3.0338 | -34.0177 | -5.3515
3.0399 | -34.0183 | -5.3652
3.046 | -34.0186 | -5.3781
3.0521 | -34.0192 | -5.3905
3.0582 | -34.0194 | -5.4042
3.0643 | -34.02 | -5.4172
3.0704 | -34.0202 | -5.4305
3.0765 | -34.0206 | -5.4426
3.0827 | -34.0209 | -5.4556
3.0888 | -34.0212 | -5.4684
3.0949 | -34.0215 | -5.4812
3.101 | -34.022 | -5.4953
3.1071 | -34.0222 | -5.5093
3.1132 | -34.0227 | -5.5232
3.1193 | -34.0231 | -5.5362
3.1254 | -34.0235 | -5.5483
3.1315 | -34.0239 | -5.5625
3.1376 | -34.0241 | -5.5753
3.1437 | -34.0244 | -5.5888
3.1498 | -34.0248 | -5.6029
3.1559 | -34.0251 | -5.616
3.162 | -34.0251 | -5.63
3.1681 | -34.0252 | -5.6441
3.1742 | -34.0255 | -5.6577
3.1803 | -34.0259 | -5.6708
3.1864 | -34.0262 | -5.683
3.1925 | -34.0265 | -5.6966
3.1986 | -34.0266 | -5.7092
3.2047 | -34.0267 | -5.7232
3.2108 | -34.0269 | -5.7348
3.2169 | -34.0273 | -5.7484
3.223 | -34.0272 | -5.7617
3.2292 | -34.0272 | -5.7755
3.2353 | -34.0271 | -5.7849
3.2414 | -34.0272 | -5.7978
3.2475 | -34.0272 | -5.8087
3.2536 | -34.0274 | -5.8223
3.2597 | -34.0276 | -5.8334
3.2658 | -34.0278 | -5.844
3.2719 | -34.0279 | -5.8549
3.278 | -34.0279 | -5.867
3.2841 | -34.028 | -5.8788
3.2902 | -34.0281 | -5.8912
3.2963 | -34.0282 | -5.9027
3.3024 | -34.0281 | -5.9127
3.3085 | -34.0283 | -5.9252
3.3146 | -34.0282 | -5.9359
3.3207 | -34.0284 | -5.948
3.3268 | -34.0285 | -5.9601
3.3329 | -34.0285 | -5.9697
3.339 | -34.0286 | -5.9792
3.3451 | -34.0286 | -5.9904
3.3512 | -34.0287 | -5.9997
3.3573 | -34.0289 | -6.0109
3.3634 | -34.0289 | -6.0223
3.3696 | -34.0289 | -6.034
3.3757 | -34.0288 | -6.0437
3.3818 | -34.0291 | -6.0552
3.3879 | -34.029 | -6.0655
3.394 | -34.0289 | -6.0758
3.4001 | -34.029 | -6.0865
3.4062 | -34.0289 | -6.0965
3.4123 | -34.029 | -6.1061
3.4184 | -34.029 | -6.117
3.4245 | -34.0291 | -6.1274
3.4306 | -34.029 | -6.1369
3.4367 | -34.0291 | -6.1465
3.4428 | -34.0294 | -6.1555
3.4489 | -34.0293 | -6.1664
3.455 | -34.0292 | -6.1767
3.4611 | -34.0292 | -6.1877
3.4672 | -34.0291 | -6.1988
3.4733 | -34.0292 | -6.2088
3.4794 | -34.0292 | -6.2191
3.4855 | -34.0294 | -6.2286
3.4916 | -34.0295 | -6.2381
3.4977 | -34.0294 | -6.2489
3.5038 | -34.0295 | -6.2587
3.5099 | -34.0296 | -6.2681
3.5161 | -34.0297 | -6.2774
3.5222 | -34.0298 | -6.2873
3.5283 | -34.0297 | -6.2972
3.5344 | -34.0297 | -6.3078
3.5405 | -34.0298 | -6.3175
3.5466 | -34.0299 | -6.3271
3.5527 | -34.0299 | -6.3376
3.5588 | -34.03 | -6.3472
3.5649 | -34.03 | -6.3571
3.571 | -34.0301 | -6.3672
3.5771 | -34.0302 | -6.3778
3.5832 | -34.0303 | -6.3872
3.5893 | -34.0305 | -6.3968
3.5954 | -34.0306 | -6.4072
3.6015 | -34.0307 | -6.4173
3.6076 | -34.0309 | -6.4268
3.6137 | -34.0309 | -6.4365
3.6198 | -34.0313 | -6.4453
3.6259 | -34.0313 | -6.4559
3.632 | -34.0315 | -6.4663
3.6381 | -34.0317 | -6.476
3.6442 | -34.0319 | -6.4862
3.6503 | -34.0321 | -6.4951
3.6565 | -34.0324 | -6.5067
3.6626 | -34.0325 | -6.5172
3.6687 | -34.0327 | -6.5276
3.6748 | -34.033 | -6.5374
3.6809 | -34.0334 | -6.5477
3.687 | -34.0336 | -6.5587
3.6931 | -34.0339 | -6.5685
3.6992 | -34.0341 | -6.5789
3.7053 | -34.0343 | -6.5888
3.7114 | -34.0345 | -6.5999
3.7175 | -34.0348 | -6.6096
3.7236 | -34.035 | -6.6182
3.7297 | -34.0351 | -6.6288
3.7358 | -34.0353 | -6.6403
3.7419 | -34.0355 | -6.6518
3.748 | -34.0357 | -6.6623
3.7541 | -34.0358 | -6.6731
3.7602 | -34.0361 | -6.6838
3.7663 | -34.0363 | -6.6938
3.7724 | -34.0366 | -6.7046
3.7785 | -34.0369 | -6.7149
3.7846 | -34.0371 | -6.7253
3.7907 | -34.0375 | -6.7361
3.7969 | -34.0377 | -6.7475
3.803 | -34.0381 | -6.7573
3.8091 | -34.0384 | -6.7675
3.8152 | -34.0386 | -6.7785
3.8213 | -34.0388 | -6.7899
3.8274 | -34.0392 | -6.8002
3.8335 | -34.0395 | -6.8117
3.8396 | -34.04 | -6.823
3.8457 | -34.0403 | -6.8331
3.8518 | -34.0408 | -6.8438
3.8579 | -34.0412 | -6.8542
3.864 | -34.0415 | -6.8647
3.8701 | -34.0419 | -6.8756
3.8762 | -34.0422 | -6.8866
3.8823 | -34.0424 | -6.897
3.8884 | -34.0429 | -6.9072
3.8945 | -34.0433 | -6.9184
3.9006 | -34.0437 | -6.9298
3.9067 | -34.044 | -6.9415
3.9128 | -34.0444 | -6.9526
3.9189 | -34.0448 | -6.963
3.925 | -34.045 | -6.9738
3.9311 | -34.0454 | -6.9836
3.9372 | -34.0458 | -6.995
3.9434 | -34.0461 | -7.0053
3.9495 | -34.0465 | -7.0167
3.9556 | -34.0469 | -7.0278
3.9617 | -34.0473 | -7.0385
3.9678 | -34.0476 | -7.049
3.9739 | -34.048 | -7.0601
3.98 | -34.0484 | -7.0717
3.9861 | -34.0486 | -7.082
3.9922 | -34.0488 | -7.0935
3.9983 | -34.0492 | -7.1034
4.0044 | -34.0495 | -7.1129
4.0105 | -34.0499 | -7.1237
4.0166 | -34.0504 | -7.1346
4.0227 | -34.0508 | -7.1447
4.0288 | -34.051 | -7.156
4.0349 | -34.0513 | -7.1657
4.041 | -34.0516 | -7.1765
4.0471 | -34.0518 | -7.1866
4.0532 | -34.0523 | -7.1976
4.0593 | -34.0525 | -7.209
4.0654 | -34.0528 | -7.2196
4.0715 | -34.0532 | -7.2308
4.0776 | -34.0535 | -7.241
4.0838 | -34.0538 | -7.2522
4.0899 | -34.0542 | -7.2633
4.096 | -34.0545 | -7.2736
4.1021 | -34.0547 | -7.2842
4.1082 | -34.0551 | -7.2943
4.1143 | -34.0554 | -7.3041
4.1204 | -34.0558 | -7.3147
4.1265 | -34.0562 | -7.3249
4.1326 | -34.0566 | -7.335
4.1387 | -34.057 | -7.3461
4.1448 | -34.0573 | -7.3569
4.1509 | -34.0575 | -7.3675
4.157 | -34.0578 | -7.3785
4.1631 | -34.058 | -7.3896
4.1692 | -34.0584 | -7.4001
4.1753 | -34.0587 | -7.4108
4.1814 | -34.059 | -7.4204
4.1875 | -34.0593 | -7.4317
4.1936 | -34.0595 | -7.4428
4.1997 | -34.0598 | -7.4528
4.2058 | -34.0601 | -7.4642
4.2119 | -34.0603 | -7.4737
4.218 | -34.0606 | -7.4837
4.2241 | -34.0609 | -7.4931
4.2303 | -34.0613 | -7.5042
4.2364 | -34.0616 | -7.5154
4.2425 | -34.0619 | -7.5257
4.2486 | -34.0621 | -7.5367
4.2547 | -34.0622 | -7.5458
4.2608 | -34.0625 | -7.5556
4.2669 | -34.0627 | -7.5658
4.273 | -34.063 | -7.5764
4.2791 | -34.0632 | -7.5869
4.2852 | -34.0634 | -7.597
4.2913 | -34.0635 | -7.6061
4.2974 | -34.0637 | -7.6168
4.3035 | -34.064 | -7.6275
4.3096 | -34.0641 | -7.6377
4.3157 | -34.0643 | -7.6475
4.3218 | -34.0645 | -7.6573
4.3279 | -34.0648 | -7.6675
4.334 | -34.0649 | -7.6772
4.3401 | -34.0651 | -7.6869
4.3462 | -34.0654 | -7.6975
4.3523 | -34.0657 | -7.707
4.3584 | -34.0658 | -7.7153
4.3645 | -34.066 | -7.7248
4.3707 | -34.0662 | -7.7339
4.3768 | -34.0664 | -7.7441
4.3829 | -34.0665 | -7.7542
4.389 | -34.0668 | -7.7637
4.3951 | -34.067 | -7.7731
4.4012 | -34.0672 | -7.7827
4.4073 | -34.0674 | -7.7926
4.4134 | -34.0676 | -7.8026
4.4195 | -34.0678 | -7.812
4.4256 | -34.0681 | -7.8214
4.4317 | -34.0683 | -7.8307
4.4378 | -34.0684 | -7.84
4.4439 | -34.0686 | -7.8493
4.45 | -34.0687 | -7.8591
4.4561 | -34.0688 | -7.8687
4.4622 | -34.0689 | -7.8781
4.4683 | -34.069 | -7.8871
4.4744 | -34.0692 | -7.8971
4.4805 | -34.0695 | -7.906
4.4866 | -34.0697 | -7.9151
4.4927 | -34.0698 | -7.9251
4.4988 | -34.07 | -7.9343
4.5049 | -34.0702 | -7.9432
4.511 | -34.0704 | -7.952
4.5172 | -34.0707 | -7.9617
4.5233 | -34.0709 | -7.9721
4.5294 | -34.0711 | -7.9812
4.5355 | -34.0713 | -7.9904
4.5416 | -34.0715 | -7.9993
4.5477 | -34.0717 | -8.0084
4.5538 | -34.0719 | -8.0169
4.5599 | -34.0722 | -8.0262
4.566 | -34.0723 | -8.0357
4.5721 | -34.0725 | -8.0448
4.5782 | -34.0728 | -8.0544
4.5843 | -34.073 | -8.0638
4.5904 | -34.0733 | -8.0729
4.5965 | -34.0736 | -8.0814
4.6026 | -34.0738 | -8.0919
4.6087 | -34.074 | -8.1007
4.6148 | -34.0742 | -8.1111
4.6209 | -34.0746 | -8.1194
4.627 | -34.0749 | -8.1292
4.6331 | -34.0749 | -8.1379
4.6392 | -34.0752 | -8.1469
4.6453 | -34.0754 | -8.1561
4.6514 | -34.0757 | -8.1654
4.6576 | -34.0757 | -8.1758
4.6637 | -34.0759 | -8.1832
4.6698 | -34.0761 | -8.1915
4.6759 | -34.0762 | -8.2012
4.682 | -34.0764 | -8.2101
4.6881 | -34.0765 | -8.2187
4.6942 | -34.0767 | -8.2277
4.7003 | -34.077 | -8.237
4.7064 | -34.0772 | -8.246
4.7125 | -34.0775 | -8.2553
4.7186 | -34.0776 | -8.2635
4.7247 | -34.0779 | -8.2729
4.7308 | -34.0782 | -8.282
4.7369 | -34.0784 | -8.2914
4.743 | -34.0786 | -8.3016
4.7491 | -34.0787 | -8.311
4.7552 | -34.0788 | -8.3201
4.7613 | -34.079 | -8.3295
4.7674 | -34.0793 | -8.3378
4.7735 | -34.0794 | -8.3478
4.7796 | -34.0796 | -8.3574
4.7857 | -34.0797 | -8.3669
4.7918 | -34.0799 | -8.3753
4.7979 | -34.0803 | -8.3841
4.8041 | -34.0803 | -8.3932
4.8102 | -34.0807 | -8.4021
4.8163 | -34.0809 | -8.4114
4.8224 | -34.081 | -8.4198
4.8285 | -34.0813 | -8.4287
4.8346 | -34.0815 | -8.438
4.8407 | -34.0817 | -8.4471
4.8468 | -34.0819 | -8.4567
4.8529 | -34.0821 | -8.4663
4.859 | -34.0823 | -8.4756
4.8651 | -34.0825 | -8.4848
4.8712 | -34.0827 | -8.494
4.8773 | -34.0829 | -8.5034
4.8834 | -34.0832 | -8.5129
4.8895 | -34.0834 | -8.5228
4.8956 | -34.0836 | -8.5325
4.9017 | -34.0837 | -8.5418
4.9078 | -34.0839 | -8.5515
4.9139 | -34.0842 | -8.5599
4.92 | -34.0844 | -8.5687
4.9261 | -34.0847 | -8.5773
4.9322 | -34.0851 | -8.5863
4.9383 | -34.0854 | -8.5952
4.9445 | -34.0856 | -8.6051
4.9506 | -34.0858 | -8.6137
4.9567 | -34.0861 | -8.6229
4.9628 | -34.0863 | -8.632
4.9689 | -34.0865 | -8.6417
4.975 | -34.0867 | -8.6509
4.9811 | -34.0869 | -8.6605
4.9872 | -34.0872 | -8.6701
4.9933 | -34.0875 | -8.6796
4.9994 | -34.0879 | -8.6893
5.0055 | -34.0882 | -8.6987
5.0116 | -34.0885 | -8.7088
5.0177 | -34.0889 | -8.7187
5.0238 | -34.0892 | -8.7291
5.0299 | -34.0896 | -8.7399
5.036 | -34.0898 | -8.7493
5.0421 | -34.09 | -8.7588
5.0482 | -34.0903 | -8.7689
5.0543 | -34.0906 | -8.778
5.0604 | -34.091 | -8.7882
5.0665 | -34.0913 | -8.7977
5.0726 | -34.0916 | -8.8079
5.0787 | -34.0919 | -8.818
5.0848 | -34.0922 | -8.8279
5.091 | -34.0925 | -8.8376
5.0971 | -34.0928 | -8.8475
5.1032 | -34.0931 | -8.8578
5.1093 | -34.0934 | -8.8675
5.1154 | -34.0937 | -8.8774
5.1215 | -34.0939 | -8.8872
5.1276 | -34.0942 | -8.8969
5.1337 | -34.0947 | -8.9064
5.1398 | -34.0949 | -8.916
5.1459 | -34.0952 | -8.9263
5.152 | -34.0955 | -8.9372
5.1581 | -34.0957 | -8.9477
5.1642 | -34.096 | -8.9576
5.1703 | -34.0964 | -8.9674
5.1764 | -34.0966 | -8.9772
5.1825 | -34.0969 | -8.9874
5.1886 | -34.0972 | -8.997
5.1947 | -34.0975 | -9.0076
5.2008 | -34.0979 | -9.0173
5.2069 | -34.0982 | -9.0284
5.213 | -34.0985 | -9.0378
5.2191 | -34.0989 | -9.048
5.2252 | -34.0991 | -9.0581
5.2314 | -34.0994 | -9.0687
5.2375 | -34.0998 | -9.0785
5.2436 | -34.1 | -9.0885
5.2497 | -34.1004 | -9.0982
5.2558 | -34.1007 | -9.1091
5.2619 | -34.101 | -9.1198
5.268 | -34.1013 | -9.1301
5.2741 | -34.1017 | -9.1412
5.2802 | -34.102 | -9.1519
5.2863 | -34.1024 | -9.1619
5.2924 | -34.1026 | -9.173
5.2985 | -34.1031 | -9.1842
5.3046 | -34.1035 | -9.1953
5.3107 | -34.1038 | -9.206
5.3168 | -34.1043 | -9.2163
5.3229 | -34.1046 | -9.226
5.329 | -34.1049 | -9.2358
5.3351 | -34.1054 | -9.2459
5.3412 | -34.1058 | -9.2559
5.3473 | -34.1061 | -9.2666
5.3534 | -34.1065 | -9.2774
5.3595 | -34.1068 | -9.2879
5.3656 | -34.1072 | -9.2988
5.3717 | -34.1075 | -9.3092
5.3779 | -34.1078 | -9.32
5.384 | -34.1081 | -9.3306
5.3901 | -34.1085 | -9.3412
5.3962 | -34.1087 | -9.3518
5.4023 | -34.1091 | -9.3624
5.4084 | -34.1095 | -9.3727
5.4145 | -34.1098 | -9.3836
5.4206 | -34.11 | -9.3948
5.4267 | -34.1104 | -9.406
5.4328 | -34.1108 | -9.416
5.4389 | -34.1111 | -9.4261
5.445 | -34.1114 | -9.4373
5.4511 | -34.1118 | -9.4476
5.4572 | -34.1122 | -9.4586
5.4633 | -34.1125 | -9.4689
5.4694 | -34.1127 | -9.4792
5.4755 | -34.1131 | -9.49
5.4816 | -34.1134 | -9.5005
5.4877 | -34.1138 | -9.511
5.4938 | -34.1141 | -9.5213
5.4999 | -34.1144 | -9.5319
5.506 | -34.1148 | -9.5425
5.5121 | -34.1151 | -9.5534
5.5183 | -34.1154 | -9.5634
5.5244 | -34.1156 | -9.574
5.5305 | -34.116 | -9.5848
5.5366 | -34.1163 | -9.5951
5.5427 | -34.1166 | -9.6062
5.5488 | -34.1169 | -9.6171
5.5549 | -34.1172 | -9.6277
5.561 | -34.1175 | -9.6388
5.5671 | -34.118 | -9.6491
5.5732 | -34.1182 | -9.6599
5.5793 | -34.1186 | -9.6713
5.5854 | -34.119 | -9.6821
5.5915 | -34.1191 | -9.6929
5.5976 | -34.1195 | -9.7038
5.6037 | -34.1198 | -9.7144
5.6098 | -34.1201 | -9.7248
5.6159 | -34.1204 | -9.7356
5.622 | -34.1208 | -9.7459
5.6281 | -34.1211 | -9.7566
5.6342 | -34.1214 | -9.7674
5.6403 | -34.1217 | -9.7775
5.6464 | -34.1221 | -9.7881
5.6525 | -34.1223 | -9.7985
5.6586 | -34.1227 | -9.8083
5.6648 | -34.123 | -9.8195
5.6709 | -34.1233 | -9.8305
5.677 | -34.1237 | -9.8419
5.6831 | -34.124 | -9.852
5.6892 | -34.1243 | -9.8626
5.6953 | -34.1246 | -9.8732
5.7014 | -34.1249 | -9.8843
5.7075 | -34.1252 | -9.8952
5.7136 | -34.1255 | -9.9056
5.7197 | -34.1258 | -9.916
5.7258 | -34.1261 | -9.9264
5.7319 | -34.1264 | -9.9371
5.738 | -34.1267 | -9.9476
5.7441 | -34.1271 | -9.9582
5.7502 | -34.1274 | -9.9684
5.7563 | -34.1278 | -9.9787
5.7624 | -34.128 | -9.99
5.7685 | -34.1283 | -10.0004
5.7746 | -34.1285 | -10.0109
5.7807 | -34.1288 | -10.0211
5.7868 | -34.1291 | -10.0316
5.7929 | -34.1294 | -10.042
5.799 | -34.1297 | -10.0525
5.8052 | -34.1299 | -10.0631
5.8113 | -34.1302 | -10.0737
5.8174 | -34.1305 | -10.0844
5.8235 | -34.1307 | -10.0947
5.8296 | -34.131 | -10.1058
5.8357 | -34.1312 | -10.1163
5.8418 | -34.1316 | -10.1271
5.8479 | -34.1318 | -10.1379
5.854 | -34.132 | -10.1487
5.8601 | -34.1324 | -10.1591
5.8662 | -34.1326 | -10.1699
5.8723 | -34.1329 | -10.1798
5.8784 | -34.1332 | -10.1903
5.8845 | -34.1335 | -10.201
5.8906 | -34.1337 | -10.2114
5.8967 | -34.134 | -10.2215
5.9028 | -34.1344 | -10.2321
5.9089 | -34.1346 | -10.242
5.915 | -34.1349 | -10.2524
5.9211 | -34.1352 | -10.263
5.9272 | -34.1355 | -10.2733
5.9333 | -34.1358 | -10.2831
5.9394 | -34.1361 | -10.2934
5.9455 | -34.1364 | -10.3034
5.9517 | -34.1368 | -10.3139
5.9578 | -34.1371 | -10.3246
5.9639 | -34.1374 | -10.3347
5.97 | -34.1377 | -10.3452
5.9761 | -34.138 | -10.3554
5.9822 | -34.1384 | -10.3659
5.9883 | -34.1387 | -10.3761
5.9944 | -34.139 | -10.3865
6.0005 | -34.1394 | -10.3972
6.0066 | -34.1396 | -10.4072
6.0127 | -34.1399 | -10.4169
6.0188 | -34.1402 | -10.4276
6.0249 | -34.1405 | -10.4379
6.031 | -34.1408 | -10.448
6.0371 | -34.1412 | -10.4587
6.0432 | -34.1415 | -10.469
6.0493 | -34.1419 | -10.4792
6.0554 | -34.1422 | -10.4894
6.0615 | -34.1425 | -10.5003
6.0676 | -34.1428 | -10.5112
6.0737 | -34.1432 | -10.5218
6.0798 | -34.1434 | -10.532
6.0859 | -34.1437 | -10.5423
6.0921 | -34.1441 | -10.5523
6.0982 | -34.1444 | -10.5629
6.1043 | -34.1448 | -10.5731
6.1104 | -34.1452 | -10.584
6.1165 | -34.1455 | -10.5935
6.1226 | -34.1458 | -10.6033
6.1287 | -34.1462 | -10.6137
6.1348 | -34.1466 | -10.624
6.1409 | -34.1469 | -10.6342
6.147 | -34.1472 | -10.6443
6.1531 | -34.1475 | -10.6546
6.1592 | -34.1479 | -10.6649
6.1653 | -34.1482 | -10.6749
6.1714 | -34.1485 | -10.6853
6.1775 | -34.1489 | -10.6959
6.1836 | -34.1491 | -10.7059
6.1897 | -34.1495 | -10.7162
6.1958 | -34.1497 | -10.7267
6.2019 | -34.1501 | -10.7373
6.208 | -34.1504 | -10.7479
6.2141 | -34.1508 | -10.7584
6.2202 | -34.1511 | -10.7691
6.2263 | -34.1515 | -10.7793
6.2325 | -34.1519 | -10.7891
6.2386 | -34.1523 | -10.799
6.2447 | -34.1525 | -10.8095
6.2508 | -34.1528 | -10.8197
6.2569 | -34.1532 | -10.8303
6.263 | -34.1536 | -10.8402
6.2691 | -34.1539 | -10.8503
6.2752 | -34.1543 | -10.8605
6.2813 | -34.1546 | -10.8707
6.2874 | -34.1549 | -10.881
6.2935 | -34.1553 | -10.891
6.2996 | -34.1556 | -10.9013
6.3057 | -34.156 | -10.9115
6.3118 | -34.1563 | -10.9217
6.3179 | -34.1567 | -10.9319
6.324 | -34.157 | -10.9424
6.3301 | -34.1574 | -10.9524
6.3362 | -34.1578 | -10.9627
6.3423 | -34.1582 | -10.973
6.3484 | -34.1586 | -10.9834
6.3545 | -34.1589 | -10.993
6.3606 | -34.1593 | -11.0029
6.3667 | -34.1597 | -11.0129
6.3728 | -34.1601 | -11.0232
6.379 | -34.1604 | -11.0333
6.3851 | -34.1608 | -11.0436
6.3912 | -34.1611 | -11.0535
6.3973 | -34.1616 | -11.0634
6.4034 | -34.1619 | -11.0735
6.4095 | -34.1622 | -11.0839
6.4156 | -34.1625 | -11.0938
6.4217 | -34.1629 | -11.1039
6.4278 | -34.1632 | -11.114
6.4339 | -34.1637 | -11.1241
6.44 | -34.164 | -11.1339
6.4461 | -34.1643 | -11.1446
6.4522 | -34.1647 | -11.1541
6.4583 | -34.165 | -11.1643
6.4644 | -34.1653 | -11.1749
6.4705 | -34.1657 | -11.1851
6.4766 | -34.1661 | -11.1957
6.4827 | -34.1664 | -11.2063
6.4888 | -34.1667 | -11.2163
6.4949 | -34.1671 | -11.226
6.501 | -34.1674 | -11.2365
6.5071 | -34.1678 | -11.2468
6.5132 | -34.1682 | -11.2568
6.5194 | -34.1686 | -11.2668
6.5255 | -34.169 | -11.2766
6.5316 | -34.1693 | -11.2869
6.5377 | -34.1697 | -11.2973
6.5438 | -34.17 | -11.3073
6.5499 | -34.1704 | -11.3173
6.556 | -34.1708 | -11.3272
6.5621 | -34.1711 | -11.3376
6.5682 | -34.1714 | -11.3475
6.5743 | -34.1718 | -11.3584
6.5804 | -34.1721 | -11.3686
6.5865 | -34.1723 | -11.379
6.5926 | -34.1728 | -11.3891
6.5987 | -34.1731 | -11.3989
6.6048 | -34.1735 | -11.4097
6.6109 | -34.1739 | -11.4197
6.617 | -34.1742 | -11.4308
6.6231 | -34.1744 | -11.4411
6.6292 | -34.1749 | -11.4512
6.6353 | -34.1752 | -11.4623
6.6414 | -34.1757 | -11.4728
6.6475 | -34.176 | -11.4832
6.6536 | -34.1764 | -11.493
6.6597 | -34.1768 | -11.5029
6.6659 | -34.1772 | -11.5129
6.672 | -34.1776 | -11.5231
6.6781 | -34.1779 | -11.533
6.6842 | -34.1783 | -11.5434
6.6903 | -34.1787 | -11.5535
6.6964 | -34.179 | -11.5637
6.7025 | -34.1793 | -11.5738
6.7086 | -34.1797 | -11.5838
6.7147 | -34.1801 | -11.5941
6.7208 | -34.1804 | -11.6048
6.7269 | -34.1807 | -11.6146
6.733 | -34.181 | -11.624
6.7391 | -34.1815 | -11.6343
6.7452 | -34.1819 | -11.6439
6.7513 | -34.1823 | -11.6537
6.7574 | -34.1826 | -11.6641
6.7635 | -34.183 | -11.6737
6.7696 | -34.1835 | -11.6839
6.7757 | -34.1839 | -11.6944
6.7818 | -34.1842 | -11.7046
6.7879 | -34.1846 | -11.7148
6.794 | -34.185 | -11.7247
6.8001 | -34.1853 | -11.7346
6.8063 | -34.1857 | -11.7449
6.8124 | -34.186 | -11.7552
6.8185 | -34.1863 | -11.7652
6.8246 | -34.1867 | -11.7754
6.8307 | -34.187 | -11.7853
6.8368 | -34.1874 | -11.7958
6.8429 | -34.1878 | -11.8058
6.849 | -34.1882 | -11.8159
6.8551 | -34.1886 | -11.826
6.8612 | -34.1889 | -11.836
6.8673 | -34.1893 | -11.8459
6.8734 | -34.1897 | -11.8567
6.8795 | -34.1901 | -11.8671
6.8856 | -34.1904 | -11.8767
6.8917 | -34.1908 | -11.8862
6.8978 | -34.1911 | -11.8962
6.9039 | -34.1915 | -11.9061
6.91 | -34.1919 | -11.9163
6.9161 | -34.1922 | -11.9262
6.9222 | -34.1926 | -11.9363
6.9283 | -34.1929 | -11.946
6.9344 | -34.1933 | -11.9559
6.9405 | -34.1936 | -11.9658
6.9466 | -34.194 | -11.976
6.9528 | -34.1943 | -11.9861
6.9589 | -34.1947 | -11.996
6.965 | -34.1951 | -12.0061
6.9711 | -34.1955 | -12.0161
6.9772 | -34.1959 | -12.0259
6.9833 | -34.1962 | -12.0362
6.9894 | -34.1965 | -12.0466
6.9955 | -34.1969 | -12.0566
7.0016 | -34.1972 | -12.0671
7.0077 | -34.1976 | -12.0774
7.0138 | -34.198 | -12.0878
7.0199 | -34.1983 | -12.0976
7.026 | -34.1986 | -12.1075
7.0321 | -34.1989 | -12.1176
7.0382 | -34.1992 | -12.1273
7.0443 | -34.1997 | -12.1377
7.0504 | -34.2 | -12.148
7.0565 | -34.2003 | -12.1583
7.0626 | -34.2007 | -12.1681
7.0687 | -34.201 | -12.1777
7.0748 | -34.2014 | -12.1876
7.0809 | -34.2017 | -12.1977
7.087 | -34.202 | -12.2078
7.0932 | -34.2024 | -12.2178
7.0993 | -34.2027 | -12.2278
7.1054 | -34.203 | -12.2376
7.1115 | -34.2033 | -12.2475
7.1176 | -34.2036 | -12.2573
7.1237 | -34.2039 | -12.2673
7.1298 | -34.2042 | -12.2768
7.1359 | -34.2045 | -12.2865
7.142 | -34.2048 | -12.2965
7.1481 | -34.2051 | -12.3062
7.1542 | -34.2054 | -12.316
7.1603 | -34.2058 | -12.3258
7.1664 | -34.2062 | -12.3358
7.1725 | -34.2065 | -12.3455
7.1786 | -34.2069 | -12.3556
7.1847 | -34.2072 | -12.3654
7.1908 | -34.2075 | -12.3754
7.1969 | -34.2078 | -12.3853
7.203 | -34.2082 | -12.3952
7.2091 | -34.2085 | -12.4049
7.2152 | -34.2088 | -12.4148
7.2213 | -34.2091 | -12.4244
7.2274 | -34.2094 | -12.4343
7.2335 | -34.2098 | -12.4442
7.2397 | -34.2101 | -12.454
7.2458 | -34.2104 | -12.4634
7.2519 | -34.2107 | -12.4733
7.258 | -34.2111 | -12.4834
7.2641 | -34.2114 | -12.4931
7.2702 | -34.2118 | -12.5033
7.2763 | -34.2121 | -12.5129
7.2824 | -34.2125 | -12.5227
7.2885 | -34.2128 | -12.5325
7.2946 | -34.2132 | -12.5425
7.3007 | -34.2134 | -12.5524
7.3068 | -34.2137 | -12.5624
7.3129 | -34.2141 | -12.572
7.319 | -34.2144 | -12.582
7.3251 | -34.2148 | -12.5915
7.3312 | -34.2151 | -12.6009
7.3373 | -34.2155 | -12.6105
7.3434 | -34.2159 | -12.62
7.3495 | -34.2162 | -12.6293
7.3556 | -34.2165 | -12.6395
7.3617 | -34.2168 | -12.6498
7.3678 | -34.2172 | -12.6597
7.3739 | -34.2175 | -12.669
7.3801 | -34.2178 | -12.6787
7.3862 | -34.2181 | -12.6884
7.3923 | -34.2184 | -12.6981
7.3984 | -34.2187 | -12.7078
7.4045 | -34.2191 | -12.7174
7.4106 | -34.2194 | -12.727
7.4167 | -34.2197 | -12.737
7.4228 | -34.2201 | -12.7467
7.4289 | -34.2204 | -12.7561
7.435 | -34.2208 | -12.7656
7.4411 | -34.2211 | -12.775
7.4472 | -34.2214 | -12.7849
7.4533 | -34.2218 | -12.7946
7.4594 | -34.2221 | -12.8046
7.4655 | -34.2225 | -12.8143
7.4716 | -34.2228 | -12.8242
7.4777 | -34.2231 | -12.8334
7.4838 | -34.2235 | -12.8433
7.4899 | -34.2238 | -12.853
7.496 | -34.2242 | -12.8627
7.5021 | -34.2245 | -12.8723
7.5082 | -34.2248 | -12.8821
7.5143 | -34.2252 | -12.8917
7.5204 | -34.2255 | -12.9011
7.5266 | -34.2258 | -12.9108
7.5327 | -34.2261 | -12.9203
7.5388 | -34.2264 | -12.93
7.5449 | -34.2267 | -12.9399
7.551 | -34.227 | -12.9498
7.5571 | -34.2273 | -12.9593
7.5632 | -34.2277 | -12.9689
7.5693 | -34.228 | -12.9787
7.5754 | -34.2284 | -12.9884
7.5815 | -34.2287 | -12.9977
7.5876 | -34.229 | -13.0074
7.5937 | -34.2293 | -13.0169
7.5998 | -34.2296 | -13.0267
7.6059 | -34.2299 | -13.0363
7.612 | -34.2303 | -13.0459
7.6181 | -34.2306 | -13.0555
7.6242 | -34.2309 | -13.0652
7.6303 | -34.2313 | -13.0747
7.6364 | -34.2316 | -13.0845
7.6425 | -34.2319 | -13.0945
7.6486 | -34.2322 | -13.1042
7.6547 | -34.2326 | -13.114
7.6608 | -34.2329 | -13.1236
7.667 | -34.2332 | -13.1327
7.6731 | -34.2335 | -13.1425
7.6792 | -34.2338 | -13.1524
7.6853 | -34.2341 | -13.1619
7.6914 | -34.2344 | -13.1717
7.6975 | -34.2347 | -13.1815
7.7036 | -34.235 | -13.1913
7.7097 | -34.2353 | -13.2009
7.7158 | -34.2356 | -13.2104
7.7219 | -34.2359 | -13.2199
7.728 | -34.2362 | -13.2297
7.7341 | -34.2365 | -13.2394
7.7402 | -34.2369 | -13.249
7.7463 | -34.2372 | -13.2587
7.7524 | -34.2374 | -13.2686
7.7585 | -34.2377 | -13.2784
7.7646 | -34.238 | -13.2878
7.7707 | -34.2383 | -13.2973
7.7768 | -34.2385 | -13.3068
7.7829 | -34.2388 | -13.3165
7.789 | -34.2391 | -13.3259
7.7951 | -34.2394 | -13.3359
7.8012 | -34.2397 | -13.346
7.8073 | -34.24 | -13.3555
7.8135 | -34.2403 | -13.3653
7.8196 | -34.2406 | -13.3749
7.8257 | -34.2409 | -13.3849
7.8318 | -34.2412 | -13.395
7.8379 | -34.2415 | -13.4042
7.844 | -34.2417 | -13.4134
7.8501 | -34.242 | -13.4234
7.8562 | -34.2422 | -13.4337
7.8623 | -34.2426 | -13.4434
7.8684 | -34.2428 | -13.4527
7.8745 | -34.2432 | -13.4623
7.8806 | -34.2434 | -13.4715
7.8867 | -34.2438 | -13.4809
7.8928 | -34.2441 | -13.4904
7.8989 | -34.2444 | -13.4996
7.905 | -34.2448 | -13.5097
7.9111 | -34.245 | -13.5194
7.9172 | -34.2453 | -13.5288
7.9233 | -34.2456 | -13.5385
7.9294 | -34.2459 | -13.5483
7.9355 | -34.2462 | -13.5574
7.9416 | -34.2464 | -13.567
7.9477 | -34.2467 | -13.5761
7.9539 | -34.247 | -13.5854
7.96 | -34.2473 | -13.5953
7.9661 | -34.2476 | -13.605
7.9722 | -34.248 | -13.6147
7.9783 | -34.2482 | -13.6243
7.9844 | -34.2486 | -13.6343
7.9905 | -34.2488 | -13.6438
7.9966 | -34.2491 | -13.6534
8.0027 | -34.2494 | -13.6634
8.0088 | -34.2497 | -13.6726
8.0149 | -34.2499 | -13.682
8.021 | -34.2502 | -13.6916
8.0271 | -34.2505 | -13.7017
8.0332 | -34.2508 | -13.7113
8.0393 | -34.2511 | -13.7211
8.0454 | -34.2514 | -13.7309
8.0515 | -34.2517 | -13.7405
8.0576 | -34.2519 | -13.7504
8.0637 | -34.2522 | -13.7595
8.0698 | -34.2525 | -13.7687
8.0759 | -34.2528 | -13.778
8.082 | -34.2531 | -13.7878
8.0881 | -34.2534 | -13.7972
8.0942 | -34.2537 | -13.8073
8.1004 | -34.2539 | -13.8173
8.1065 | -34.2543 | -13.8269
8.1126 | -34.2546 | -13.8364
8.1187 | -34.2549 | -13.8461
8.1248 | -34.2552 | -13.856
8.1309 | -34.2555 | -13.8656
8.137 | -34.2557 | -13.8751
8.1431 | -34.256 | -13.8849
8.1492 | -34.2563 | -13.8946
8.1553 | -34.2566 | -13.9042
8.1614 | -34.2569 | -13.9141
8.1675 | -34.2572 | -13.9236
8.1736 | -34.2575 | -13.9335
8.1797 | -34.2577 | -13.9433
8.1858 | -34.258 | -13.9529
8.1919 | -34.2582 | -13.9625
8.198 | -34.2585 | -13.9718
8.2041 | -34.2589 | -13.9818
8.2102 | -34.2592 | -13.9914
8.2163 | -34.2595 | -14.001
8.2224 | -34.2598 | -14.0108
8.2285 | -34.2601 | -14.0202
8.2346 | -34.2604 | -14.0298
8.2408 | -34.2607 | -14.0395
8.2469 | -34.2609 | -14.049
8.253 | -34.2612 | -14.0586
8.2591 | -34.2615 | -14.068
8.2652 | -34.2619 | -14.0774
8.2713 | -34.2622 | -14.0871
8.2774 | -34.2625 | -14.0968
8.2835 | -34.2628 | -14.1069
8.2896 | -34.2631 | -14.116
8.2957 | -34.2634 | -14.1256
8.3018 | -34.2637 | -14.1352
8.3079 | -34.264 | -14.1448
8.314 | -34.2643 | -14.1542
8.3201 | -34.2646 | -14.1638
8.3262 | -34.2649 | -14.1734
8.3323 | -34.2652 | -14.1829
8.3384 | -34.2655 | -14.193
8.3445 | -34.2658 | -14.2022
8.3506 | -34.2661 | -14.212
8.3567 | -34.2664 | -14.2219
8.3628 | -34.2666 | -14.2314
8.3689 | -34.2669 | -14.2411
8.375 | -34.2672 | -14.2507
8.3812 | -34.2675 | -14.2603
8.3873 | -34.2678 | -14.27
8.3934 | -34.2681 | -14.2796
8.3995 | -34.2684 | -14.2893
8.4056 | -34.2687 | -14.2991
8.4117 | -34.269 | -14.309
8.4178 | -34.2693 | -14.3182
8.4239 | -34.2696 | -14.3276
8.43 | -34.2699 | -14.3372
8.4361 | -34.2702 | -14.3472
8.4422 | -34.2705 | -14.3568
8.4483 | -34.2707 | -14.3664
8.4544 | -34.2711 | -14.3763
8.4605 | -34.2713 | -14.3858
8.4666 | -34.2716 | -14.3956
8.4727 | -34.2719 | -14.4051
8.4788 | -34.2722 | -14.4145
8.4849 | -34.2725 | -14.4245
8.491 | -34.2727 | -14.4339
8.4971 | -34.273 | -14.4437
8.5032 | -34.2733 | -14.4533
8.5093 | -34.2737 | -14.4629
8.5154 | -34.2739 | -14.4723
8.5215 | -34.2742 | -14.482
8.5277 | -34.2745 | -14.4917
8.5338 | -34.2748 | -14.501
8.5399 | -34.2751 | -14.5109
8.546 | -34.2754 | -14.5202
8.5521 | -34.2757 | -14.53
8.5582 | -34.276 | -14.5398
8.5643 | -34.2763 | -14.5492
8.5704 | -34.2766 | -14.5585
8.5765 | -34.2769 | -14.5681
8.5826 | -34.2773 | -14.5776
8.5887 | -34.2775 | -14.5871
8.5948 | -34.2778 | -14.5963
8.6009 | -34.2781 | -14.6058
8.607 | -34.2784 | -14.6157
8.6131 | -34.2786 | -14.6255
8.6192 | -34.279 | -14.6348
8.6253 | -34.2793 | -14.6445
8.6314 | -34.2796 | -14.6541
8.6375 | -34.2799 | -14.6639
8.6436 | -34.2803 | -14.6735
8.6497 | -34.2806 | -14.6832
8.6558 | -34.2808 | -14.6926
8.6619 | -34.2811 | -14.7024
8.6681 | -34.2814 | -14.7119
8.6742 | -34.2817 | -14.7217
8.6803 | -34.282 | -14.7316
8.6864 | -34.2824 | -14.7415
8.6925 | -34.2827 | -14.751
8.6986 | -34.283 | -14.761
8.7047 | -34.2833 | -14.7706
8.7108 | -34.2836 | -14.7804
8.7169 | -34.2839 | -14.7898
8.723 | -34.2842 | -14.7994
8.7291 | -34.2845 | -14.8089
8.7352 | -34.2849 | -14.8188
8.7413 | -34.2851 | -14.8283
8.7474 | -34.2855 | -14.8376
8.7535 | -34.2858 | -14.8474
8.7596 | -34.2862 | -14.8575
8.7657 | -34.2864 | -14.8671
8.7718 | -34.2867 | -14.8765
8.7779 | -34.2871 | -14.8867
8.784 | -34.2873 | -14.8961
8.7901 | -34.2876 | -14.9062
8.7962 | -34.2879 | -14.9159
8.8023 | -34.2883 | -14.9252
8.8084 | -34.2886 | -14.935
8.8146 | -34.2889 | -14.9445
8.8207 | -34.2892 | -14.9542
8.8268 | -34.2896 | -14.9637
8.8329 | -34.2899 | -14.9734
8.839 | -34.2902 | -14.9831
8.8451 | -34.2905 | -14.9926
8.8512 | -34.2909 | -15.0023
8.8573 | -34.2912 | -15.0117
8.8634 | -34.2916 | -15.0212
8.8695 | -34.2919 | -15.031
8.8756 | -34.2922 | -15.0404
8.8817 | -34.2925 | -15.0502
8.8878 | -34.2927 | -15.0601
8.8939 | -34.293 | -15.0694
8.9 | -34.2934 | -15.0787
8.9061 | -34.2937 | -15.0884
8.9122 | -34.294 | -15.0979
8.9183 | -34.2943 | -15.1074
8.9244 | -34.2946 | -15.1168
8.9305 | -34.295 | -15.1265
8.9366 | -34.2954 | -15.1359
8.9427 | -34.2957 | -15.1456
8.9488 | -34.296 | -15.1552
8.955 | -34.2964 | -15.1647
8.9611 | -34.2967 | -15.1745
8.9672 | -34.2971 | -15.1845
8.9733 | -34.2974 | -15.1939
8.9794 | -34.2977 | -15.2037
8.9855 | -34.2981 | -15.2132
8.9916 | -34.2984 | -15.2227
8.9977 | -34.2987 | -15.2322
9.0038 | -34.2991 | -15.242
9.0099 | -34.2994 | -15.2517
9.016 | -34.2998 | -15.2613
9.0221 | -34.3001 | -15.2712
9.0282 | -34.3005 | -15.2809
9.0343 | -34.3008 | -15.2908
9.0404 | -34.3012 | -15.3003
9.0465 | -34.3016 | -15.3103
9.0526 | -34.3019 | -15.3198
9.0587 | -34.3023 | -15.3294
9.0648 | -34.3027 | -15.3389
9.0709 | -34.303 | -15.3483
9.077 | -34.3033 | -15.3585
9.0831 | -34.3036 | -15.3681
9.0892 | -34.3039 | -15.3777
9.0953 | -34.3042 | -15.3874
9.1015 | -34.3046 | -15.3971
9.1076 | -34.305 | -15.4068
9.1137 | -34.3054 | -15.4163
9.1198 | -34.3057 | -15.4255
9.1259 | -34.3061 | -15.4352
9.132 | -34.3065 | -15.4451
9.1381 | -34.3069 | -15.4546
9.1442 | -34.3073 | -15.4638
9.1503 | -34.3076 | -15.4735
9.1564 | -34.3079 | -15.4831
9.1625 | -34.3082 | -15.4928
9.1686 | -34.3086 | -15.5024
9.1747 | -34.309 | -15.5121
9.1808 | -34.3094 | -15.522
9.1869 | -34.3097 | -15.5319
9.193 | -34.31 | -15.5417
9.1991 | -34.3103 | -15.551
9.2052 | -34.3107 | -15.5605
9.2113 | -34.3111 | -15.5704
9.2174 | -34.3115 | -15.5806
9.2235 | -34.3119 | -15.59
9.2296 | -34.3123 | -15.5997
9.2357 | -34.3126 | -15.6094
9.2419 | -34.313 | -15.619
9.248 | -34.3134 | -15.6288
9.2541 | -34.3137 | -15.6387
9.2602 | -34.3141 | -15.6484
9.2663 | -34.3144 | -15.6579
9.2724 | -34.3148 | -15.6676
9.2785 | -34.3151 | -15.6772
9.2846 | -34.3155 | -15.6869
9.2907 | -34.3159 | -15.6965
9.2968 | -34.3162 | -15.7062
9.3029 | -34.3165 | -15.716
9.309 | -34.3169 | -15.7256
9.3151 | -34.3173 | -15.7351
9.3212 | -34.3177 | -15.7449
9.3273 | -34.3181 | -15.7544
9.3334 | -34.3185 | -15.7638
9.3395 | -34.3189 | -15.7733
9.3456 | -34.3194 | -15.7827
9.3517 | -34.3197 | -15.792
9.3578 | -34.3201 | -15.8014
9.3639 | -34.3205 | -15.8112
9.37 | -34.3209 | -15.8209
9.3761 | -34.3213 | -15.8304
9.3822 | -34.3216 | -15.8399
9.3884 | -34.322 | -15.8495
9.3945 | -34.3224 | -15.8592
9.4006 | -34.3228 | -15.8689
9.4067 | -34.3231 | -15.8786
9.4128 | -34.3235 | -15.8882
9.4189 | -34.3239 | -15.8979
9.425 | -34.3243 | -15.9074
9.4311 | -34.3247 | -15.917
9.4372 | -34.3251 | -15.9266
9.4433 | -34.3254 | -15.9364
9.4494 | -34.3258 | -15.9459
9.4555 | -34.3262 | -15.9555
9.4616 | -34.3266 | -15.965
9.4677 | -34.3271 | -15.9745
9.4738 | -34.3274 | -15.9845
9.4799 | -34.3277 | -15.994
9.486 | -34.3281 | -16.0039
9.4921 | -34.3285 | -16.0139
9.4982 | -34.3289 | -16.0231
9.5043 | -34.3293 | -16.0324
9.5104 | -34.3297 | -16.0418
9.5165 | -34.3301 | -16.0514
9.5226 | -34.3305 | -16.0612
9.5288 | -34.3309 | -16.0707
9.5349 | -34.3313 | -16.0803
9.541 | -34.3317 | -16.0898
9.5471 | -34.332 | -16.0995
9.5532 | -34.3324 | -16.1091
9.5593 | -34.3329 | -16.1187
9.5654 | -34.3333 | -16.1283
9.5715 | -34.3336 | -16.1382
9.5776 | -34.334 | -16.148
9.5837 | -34.3344 | -16.1573
9.5898 | -34.3348 | -16.1671
9.5959 | -34.3351 | -16.1766
9.602 | -34.3355 | -16.1863
9.6081 | -34.3359 | -16.196
9.6142 | -34.3363 | -16.2055
9.6203 | -34.3366 | -16.2152
9.6264 | -34.337 | -16.2244
9.6325 | -34.3373 | -16.2342
9.6386 | -34.3377 | -16.2437
9.6447 | -34.3381 | -16.2532
9.6508 | -34.3385 | -16.2629
9.6569 | -34.3389 | -16.2724
9.663 | -34.3393 | -16.282
9.6691 | -34.3397 | -16.2916
9.6753 | -34.3402 | -16.3012
9.6814 | -34.3406 | -16.3105
9.6875 | -34.341 | -16.3199
9.6936 | -34.3414 | -16.3293
9.6997 | -34.3418 | -16.3391
9.7058 | -34.3422 | -16.3486
9.7119 | -34.3426 | -16.3581
9.718 | -34.343 | -16.368
9.7241 | -34.3433 | -16.3776
9.7302 | -34.3437 | -16.3872
9.7363 | -34.3441 | -16.3966
9.7424 | -34.3446 | -16.406
9.7485 | -34.3449 | -16.4155
9.7546 | -34.3454 | -16.4248
9.7607 | -34.3457 | -16.4345
9.7668 | -34.3461 | -16.4445
9.7729 | -34.3466 | -16.4538
9.779 | -34.3471 | -16.4634
9.7851 | -34.3474 | -16.4731
9.7912 | -34.3478 | -16.4832
9.7973 | -34.3481 | -16.493
9.8034 | -34.3486 | -16.5027
9.8095 | -34.3489 | -16.5122
9.8157 | -34.3493 | -16.5218
9.8218 | -34.3497 | -16.531
9.8279 | -34.3501 | -16.5406
9.834 | -34.3506 | -16.55
9.8401 | -34.351 | -16.5595
9.8462 | -34.3514 | -16.5687
9.8523 | -34.3518 | -16.5781
9.8584 | -34.3522 | -16.5879
9.8645 | -34.3526 | -16.5975
9.8706 | -34.353 | -16.6069
9.8767 | -34.3533 | -16.6161
9.8828 | -34.3538 | -16.6253
9.8889 | -34.3542 | -16.6351
9.895 | -34.3545 | -16.6443
9.9011 | -34.3549 | -16.6533
9.9072 | -34.3554 | -16.6625
9.9133 | -34.3558 | -16.6718
9.9194 | -34.3562 | -16.681
9.9255 | -34.3566 | -16.6906
9.9316 | -34.357 | -16.6999
9.9377 | -34.3574 | -16.7093
9.9438 | -34.3579 | -16.7186
9.9499 | -34.3583 | -16.7282
9.956 | -34.3587 | -16.7375
9.9622 | -34.3591 | -16.7469
9.9683 | -34.3595 | -16.7564
9.9744 | -34.3599 | -16.7658
9.9805 | -34.3603 | -16.7753
9.9866 | -34.3607 | -16.7844
9.9927 | -34.3611 | -16.794
9.9988 | -34.3615 | -16.8034
10.0049 | -34.3619 | -16.8128
10.011 | -34.3624 | -16.8221
10.0171 | -34.3627 | -16.8315
10.0232 | -34.3632 | -16.8406
10.0293 | -34.3636 | -16.8499
10.0354 | -34.3641 | -16.8591
10.0415 | -34.3645 | -16.8684
10.0476 | -34.3649 | -16.8782
10.0537 | -34.3652 | -16.8878
10.0598 | -34.3657 | -16.8973
10.0659 | -34.366 | -16.9067
10.072 | -34.3665 | -16.916
10.0781 | -34.3669 | -16.9254
10.0842 | -34.3673 | -16.935
10.0903 | -34.3677 | -16.9447
10.0964 | -34.3682 | -16.9538
10.1026 | -34.3686 | -16.9632
10.1087 | -34.3691 | -16.9726
10.1148 | -34.3695 | -16.9821
10.1209 | -34.3699 | -16.9912
10.127 | -34.3703 | -17.0004
10.1331 | -34.3707 | -17.0097
10.1392 | -34.3711 | -17.019
10.1453 | -34.3715 | -17.0284
10.1514 | -34.3719 | -17.0375
10.1575 | -34.3723 | -17.0468
10.1636 | -34.3728 | -17.0554
10.1697 | -34.3732 | -17.0648
10.1758 | -34.3736 | -17.0739
10.1819 | -34.374 | -17.0834
10.188 | -34.3744 | -17.0924
10.1941 | -34.3748 | -17.1019
10.2002 | -34.3752 | -17.1107
10.2063 | -34.3756 | -17.1202
10.2124 | -34.376 | -17.1296
10.2185 | -34.3764 | -17.1388
10.2246 | -34.3768 | -17.1481
10.2307 | -34.3772 | -17.1576
10.2368 | -34.3776 | -17.1659
10.2429 | -34.3781 | -17.175
10.2491 | -34.3785 | -17.1841
10.2552 | -34.3789 | -17.1938
10.2613 | -34.3794 | -17.2033
10.2674 | -34.3797 | -17.2127
10.2735 | -34.3801 | -17.2223
10.2796 | -34.3805 | -17.2316
10.2857 | -34.3809 | -17.2409
10.2918 | -34.3813 | -17.2504
10.2979 | -34.3817 | -17.2599
10.304 | -34.3822 | -17.2694
10.3101 | -34.3826 | -17.2787
10.3162 | -34.383 | -17.288
10.3223 | -34.3834 | -17.2973
10.3284 | -34.3839 | -17.3065
10.3345 | -34.3843 | -17.3157
10.3406 | -34.3847 | -17.3248
10.3467 | -34.3851 | -17.334
10.3528 | -34.3855 | -17.343
10.3589 | -34.3858 | -17.3524
10.365 | -34.3863 | -17.362
10.3711 | -34.3866 | -17.3716
10.3772 | -34.3869 | -17.3808
10.3833 | -34.3874 | -17.39
10.3895 | -34.3878 | -17.3991
10.3956 | -34.3882 | -17.4082
10.4017 | -34.3886 | -17.4173
10.4078 | -34.389 | -17.4266
10.4139 | -34.3894 | -17.4357
10.42 | -34.3898 | -17.445
10.4261 | -34.3903 | -17.4542
10.4322 | -34.3907 | -17.4632
10.4383 | -34.3911 | -17.4724
10.4444 | -34.3915 | -17.4812
10.4505 | -34.3919 | -17.4904
10.4566 | -34.3923 | -17.4995
10.4627 | -34.3927 | -17.5087
10.4688 | -34.3932 | -17.5174
10.4749 | -34.3936 | -17.5265
10.481 | -34.394 | -17.5357
10.4871 | -34.3944 | -17.545
10.4932 | -34.3949 | -17.5542
10.4993 | -34.3953 | -17.5632
10.5054 | -34.3957 | -17.5723
10.5115 | -34.3961 | -17.5815
10.5176 | -34.3966 | -17.5906
10.5237 | -34.397 | -17.5997
10.5298 | -34.3974 | -17.6086
10.536 | -34.3978 | -17.618
10.5421 | -34.3982 | -17.6267
10.5482 | -34.3986 | -17.6362
10.5543 | -34.3991 | -17.6454
10.5604 | -34.3995 | -17.6541
10.5665 | -34.3999 | -17.6631
10.5726 | -34.4003 | -17.6719
10.5787 | -34.4007 | -17.6812
10.5848 | -34.4012 | -17.6901
10.5909 | -34.4016 | -17.6993
10.597 | -34.4019 | -17.7085
10.6031 | -34.4024 | -17.7179
10.6092 | -34.4028 | -17.7273
10.6153 | -34.4032 | -17.7366
10.6214 | -34.4036 | -17.7455
10.6275 | -34.404 | -17.7544
10.6336 | -34.4044 | -17.7638
10.6397 | -34.4048 | -17.7728
10.6458 | -34.4052 | -17.7815
10.6519 | -34.4056 | -17.7904
10.658 | -34.4059 | -17.7993
10.6641 | -34.4064 | -17.8083
10.6702 | -34.4068 | -17.8176
10.6764 | -34.4073 | -17.8266
10.6825 | -34.4077 | -17.8359
10.6886 | -34.4081 | -17.8452
10.6947 | -34.4085 | -17.8539
10.7008 | -34.4089 | -17.863
10.7069 | -34.4093 | -17.8721
10.713 | -34.4098 | -17.8811
10.7191 | -34.4102 | -17.8903
10.7252 | -34.4106 | -17.899
10.7313 | -34.4111 | -17.9081
10.7374 | -34.4115 | -17.9169
10.7435 | -34.4119 | -17.9262
10.7496 | -34.4123 | -17.9351
10.7557 | -34.4127 | -17.9443
10.7618 | -34.4132 | -17.9533
10.7679 | -34.4136 | -17.9623
10.774 | -34.4139 | -17.9713
10.7801 | -34.4143 | -17.9808
10.7862 | -34.4148 | -17.9897
10.7923 | -34.4152 | -17.999
10.7984 | -34.4156 | -18.008
10.8045 | -34.416 | -18.0173
10.8106 | -34.4164 | -18.0263
10.8168 | -34.4168 | -18.0353
10.8229 | -34.4172 | -18.0444
10.829 | -34.4177 | -18.0534
10.8351 | -34.4181 | -18.0624
10.8412 | -34.4184 | -18.0714
10.8473 | -34.4188 | -18.0801
10.8534 | -34.4193 | -18.0893
10.8595 | -34.4197 | -18.0981
10.8656 | -34.4201 | -18.1069
10.8717 | -34.4206 | -18.1156
10.8778 | -34.421 | -18.1245
10.8839 | -34.4214 | -18.1333
10.89 | -34.4218 | -18.142
10.8961 | -34.4222 | -18.1511
10.9022 | -34.4227 | -18.16
10.9083 | -34.4231 | -18.1691
10.9144 | -34.4236 | -18.178
10.9205 | -34.424 | -18.187
10.9266 | -34.4245 | -18.1962
10.9327 | -34.4249 | -18.2051
10.9388 | -34.4253 | -18.2142
10.9449 | -34.4257 | -18.2232
10.951 | -34.4261 | -18.232
10.9571 | -34.4265 | -18.241
10.9633 | -34.4269 | -18.2503
10.9694 | -34.4273 | -18.2597
10.9755 | -34.4278 | -18.2687
10.9816 | -34.4282 | -18.2777
10.9877 | -34.4286 | -18.2868
10.9938 | -34.429 | -18.296
10.9999 | -34.4295 | -18.3049
11.006 | -34.4299 | -18.3136
11.0121 | -34.4303 | -18.3229
11.0182 | -34.4308 | -18.3316
11.0243 | -34.4312 | -18.3406
11.0304 | -34.4316 | -18.3495
11.0365 | -34.432 | -18.3584
11.0426 | -34.4324 | -18.3672
11.0487 | -34.4329 | -18.3762
11.0548 | -34.4333 | -18.3851
11.0609 | -34.4337 | -18.3943
11.067 | -34.4341 | -18.4035
11.0731 | -34.4346 | -18.4124
11.0792 | -34.435 | -18.4213
11.0853 | -34.4354 | -18.43
11.0914 | -34.4358 | -18.4389
11.0975 | -34.4363 | -18.4481
11.1037 | -34.4367 | -18.4566
11.1098 | -34.4371 | -18.4659
11.1159 | -34.4375 | -18.4749
11.122 | -34.438 | -18.4842
11.1281 | -34.4383 | -18.4929
11.1342 | -34.4387 | -18.5016
11.1403 | -34.4392 | -18.5109
11.1464 | -34.4396 | -18.5198
11.1525 | -34.44 | -18.5288
11.1586 | -34.4404 | -18.5383
11.1647 | -34.4409 | -18.5471
11.1708 | -34.4413 | -18.5558
11.1769 | -34.4417 | -18.5643
11.183 | -34.4421 | -18.5731
11.1891 | -34.4425 | -18.5824
11.1952 | -34.4429 | -18.5917
11.2013 | -34.4433 | -18.6006
11.2074 | -34.4438 | -18.6092
11.2135 | -34.4441 | -18.6182
11.2196 | -34.4445 | -18.6271
11.2257 | -34.4448 | -18.6358
11.2318 | -34.4452 | -18.6449
11.2379 | -34.4456 | -18.6537
11.244 | -34.446 | -18.6627
11.2502 | -34.4464 | -18.6714
11.2563 | -34.4469 | -18.6803
11.2624 | -34.4473 | -18.6892
11.2685 | -34.4477 | -18.6979
11.2746 | -34.4481 | -18.7064
11.2807 | -34.4485 | -18.7158
11.2868 | -34.4488 | -18.725
11.2929 | -34.4493 | -18.7336
11.299 | -34.4497 | -18.7427
11.3051 | -34.4501 | -18.7516
11.3112 | -34.4505 | -18.7606
11.3173 | -34.4509 | -18.7694
11.3234 | -34.4513 | -18.7784
11.3295 | -34.4517 | -18.7871
11.3356 | -34.4522 | -18.7959
11.3417 | -34.4526 | -18.8049
11.3478 | -34.453 | -18.8138
11.3539 | -34.4533 | -18.8226
11.36 | -34.4538 | -18.8315
11.3661 | -34.4542 | -18.8403
11.3722 | -34.4546 | -18.849
11.3783 | -34.455 | -18.8579
11.3844 | -34.4554 | -18.8668
11.3906 | -34.4558 | -18.8759
11.3967 | -34.4562 | -18.8847
11.4028 | -34.4566 | -18.8937
11.4089 | -34.457 | -18.9023
11.415 | -34.4574 | -18.9111
11.4211 | -34.4578 | -18.9198
11.4272 | -34.4582 | -18.9288
11.4333 | -34.4586 | -18.9377
11.4394 | -34.459 | -18.9468
11.4455 | -34.4593 | -18.9558
11.4516 | -34.4598 | -18.9645
11.4577 | -34.4602 | -18.9736
11.4638 | -34.4607 | -18.982
11.4699 | -34.461 | -18.9912
11.476 | -34.4614 | -19.0005
11.4821 | -34.4618 | -19.0091
11.4882 | -34.4623 | -19.0183
11.4943 | -34.4627 | -19.0269
11.5004 | -34.4631 | -19.0359
11.5065 | -34.4636 | -19.0446
11.5126 | -34.4639 | -19.0532
11.5187 | -34.4644 | -19.062
11.5248 | -34.4648 | -19.0709
11.5309 | -34.4652 | -19.0797
11.5371 | -34.4656 | -19.0885
11.5432 | -34.4659 | -19.0975
11.5493 | -34.4663 | -19.1065
11.5554 | -34.4667 | -19.1151
11.5615 | -34.4671 | -19.1243
11.5676 | -34.4676 | -19.1334
11.5737 | -34.468 | -19.1419
11.5798 | -34.4685 | -19.1506
11.5859 | -34.4689 | -19.1593
11.592 | -34.4693 | -19.1682
11.5981 | -34.4697 | -19.1769
11.6042 | -34.4701 | -19.1858
11.6103 | -34.4705 | -19.1945
11.6164 | -34.4709 | -19.2032
11.6225 | -34.4713 | -19.212
11.6286 | -34.4717 | -19.2209
11.6347 | -34.4721 | -19.23
11.6408 | -34.4726 | -19.2389
11.6469 | -34.473 | -19.2479
11.653 | -34.4734 | -19.2565
11.6591 | -34.4739 | -19.2654
11.6652 | -34.4743 | -19.2742
11.6713 | -34.4747 | -19.283
11.6775 | -34.4751 | -19.2916
11.6836 | -34.4755 | -19.3005
11.6897 | -34.476 | -19.3091
11.6958 | -34.4764 | -19.3183
11.7019 | -34.4768 | -19.3273
11.708 | -34.4772 | -19.3363
11.7141 | -34.4777 | -19.3452
11.7202 | -34.4781 | -19.3538
11.7263 | -34.4785 | -19.3627
11.7324 | -34.4789 | -19.3717
11.7385 | -34.4793 | -19.3808
11.7446 | -34.4797 | -19.3898
11.7507 | -34.4802 | -19.3985
11.7568 | -34.4806 | -19.4075
11.7629 | -34.481 | -19.4163
11.769 | -34.4814 | -19.4251
11.7751 | -34.4819 | -19.4336
11.7812 | -34.4823 | -19.4424
11.7873 | -34.4827 | -19.4515
11.7934 | -34.4831 | -19.4603
11.7995 | -34.4835 | -19.4688
11.8056 | -34.4839 | -19.4779
11.8117 | -34.4843 | -19.4868
11.8178 | -34.4847 | -19.4957
11.824 | -34.4852 | -19.5043
11.8301 | -34.4857 | -19.5127
11.8362 | -34.4861 | -19.5212
11.8423 | -34.4865 | -19.53
11.8484 | -34.4869 | -19.5388
11.8545 | -34.4873 | -19.5479
11.8606 | -34.4877 | -19.5567
11.8667 | -34.4882 | -19.5656
11.8728 | -34.4886 | -19.5744
11.8789 | -34.489 | -19.5836
11.885 | -34.4893 | -19.5923
11.8911 | -34.4898 | -19.6013
11.8972 | -34.4902 | -19.6098
11.9033 | -34.4906 | -19.6181
11.9094 | -34.491 | -19.6266
11.9155 | -34.4914 | -19.6354
11.9216 | -34.4919 | -19.6441
11.9277 | -34.4923 | -19.6527
11.9338 | -34.4927 | -19.6615
11.9399 | -34.4931 | -19.6706
11.946 | -34.4936 | -19.6794
11.9521 | -34.494 | -19.6881
11.9582 | -34.4944 | -19.6967
11.9644 | -34.4949 | -19.7053
11.9705 | -34.4953 | -19.7143
11.9766 | -34.4957 | -19.723
11.9827 | -34.4961 | -19.7317
11.9888 | -34.4965 | -19.7403
11.9949 | -34.4969 | -19.749
12.001 | -34.4974 | -19.7578
12.0071 | -34.4978 | -19.7664
12.0132 | -34.4982 | -19.7748
12.0193 | -34.4987 | -19.7839
12.0254 | -34.4991 | -19.7924
12.0315 | -34.4995 | -19.8011
12.0376 | -34.4999 | -19.8095
12.0437 | -34.5003 | -19.8181
12.0498 | -34.5007 | -19.8271
12.0559 | -34.5012 | -19.8359
12.062 | -34.5016 | -19.8446
12.0681 | -34.502 | -19.8532
12.0742 | -34.5025 | -19.8621
12.0803 | -34.5029 | -19.8708
12.0864 | -34.5033 | -19.8795
12.0925 | -34.5038 | -19.8883
12.0986 | -34.5042 | -19.8968
12.1047 | -34.5047 | -19.9055
12.1109 | -34.5051 | -19.9144
12.117 | -34.5056 | -19.923
12.1231 | -34.506 | -19.9317
12.1292 | -34.5064 | -19.9406
12.1353 | -34.5069 | -19.9491
12.1414 | -34.5073 | -19.958
12.1475 | -34.5077 | -19.9668
12.1536 | -34.5081 | -19.9756
12.1597 | -34.5085 | -19.9844
12.1658 | -34.509 | -19.9931
12.1719 | -34.5095 | -20.002
12.178 | -34.5099 | -20.0109
12.1841 | -34.5103 | -20.0198
12.1902 | -34.5108 | -20.0285
12.1963 | -34.5112 | -20.0372
12.2024 | -34.5117 | -20.0459
12.2085 | -34.5121 | -20.0546
12.2146 | -34.5125 | -20.0632
12.2207 | -34.5129 | -20.0719
12.2268 | -34.5133 | -20.0806
12.2329 | -34.5138 | -20.0893
12.239 | -34.5142 | -20.0978
12.2451 | -34.5146 | -20.1062
12.2513 | -34.5151 | -20.115
12.2574 | -34.5155 | -20.124
12.2635 | -34.516 | -20.1323
12.2696 | -34.5164 | -20.1408
12.2757 | -34.5168 | -20.1492
12.2818 | -34.5172 | -20.1576
12.2879 | -34.5176 | -20.1662
12.294 | -34.518 | -20.1746
12.3001 | -34.5185 | -20.1835
12.3062 | -34.519 | -20.1922
12.3123 | -34.5194 | -20.201
12.3184 | -34.5198 | -20.2095
12.3245 | -34.5203 | -20.2183
12.3306 | -34.5207 | -20.2267
12.3367 | -34.5211 | -20.2352
12.3428 | -34.5216 | -20.2437
12.3489 | -34.522 | -20.2519
12.355 | -34.5224 | -20.2605
12.3611 | -34.5229 | -20.2692
12.3672 | -34.5233 | -20.2777
12.3733 | -34.5237 | -20.2867
12.3794 | -34.5241 | -20.2949
12.3855 | -34.5246 | -20.3034
12.3916 | -34.525 | -20.3122
12.3978 | -34.5254 | -20.3211
12.4039 | -34.5258 | -20.3297
12.41 | -34.5263 | -20.3385
12.4161 | -34.5267 | -20.347
12.4222 | -34.5272 | -20.3555
12.4283 | -34.5276 | -20.364
12.4344 | -34.5281 | -20.3727
12.4405 | -34.5285 | -20.3814
12.4466 | -34.5289 | -20.3899
12.4527 | -34.5294 | -20.3986
12.4588 | -34.5298 | -20.4072
12.4649 | -34.5302 | -20.4161
12.471 | -34.5307 | -20.4252
12.4771 | -34.5311 | -20.4332
12.4832 | -34.5315 | -20.4418
12.4893 | -34.532 | -20.4502
12.4954 | -34.5325 | -20.4587
12.5015 | -34.533 | -20.467
12.5076 | -34.5334 | -20.4757
12.5137 | -34.5338 | -20.4848
12.5198 | -34.5342 | -20.4931
12.5259 | -34.5347 | -20.5016
12.532 | -34.5351 | -20.5104
12.5382 | -34.5356 | -20.5188
12.5443 | -34.536 | -20.5275
12.5504 | -34.5364 | -20.5362
12.5565 | -34.5369 | -20.5446
12.5626 | -34.5374 | -20.5534
12.5687 | -34.5378 | -20.5622
12.5748 | -34.5382 | -20.5707
12.5809 | -34.5386 | -20.5789
12.587 | -34.539 | -20.5878
12.5931 | -34.5395 | -20.5965
12.5992 | -34.5399 | -20.6054
12.6053 | -34.5404 | -20.6138
12.6114 | -34.5409 | -20.6225
12.6175 | -34.5413 | -20.6312
12.6236 | -34.5417 | -20.6395
12.6297 | -34.5422 | -20.648
12.6358 | -34.5427 | -20.6564
12.6419 | -34.5431 | -20.6652
12.648 | -34.5436 | -20.674
12.6541 | -34.544 | -20.6825
12.6602 | -34.5445 | -20.6907
12.6663 | -34.5449 | -20.6991
12.6724 | -34.5453 | -20.7077
12.6785 | -34.5458 | -20.7166
12.6847 | -34.5462 | -20.7252
12.6908 | -34.5467 | -20.7337
12.6969 | -34.5472 | -20.7423
12.703 | -34.5476 | -20.7509
12.7091 | -34.548 | -20.7595
12.7152 | -34.5485 | -20.768
12.7213 | -34.5489 | -20.7767
12.7274 | -34.5493 | -20.7853
12.7335 | -34.5498 | -20.7934
12.7396 | -34.5502 | -20.8016
12.7457 | -34.5507 | -20.8102
12.7518 | -34.5511 | -20.8187
12.7579 | -34.5515 | -20.8276
12.764 | -34.5519 | -20.8362
12.7701 | -34.5523 | -20.8445
12.7762 | -34.5527 | -20.8531
12.7823 | -34.5532 | -20.8616
12.7884 | -34.5536 | -20.8701
12.7945 | -34.554 | -20.8783
12.8006 | -34.5545 | -20.8863
12.8067 | -34.5549 | -20.895
12.8128 | -34.5554 | -20.9032
12.8189 | -34.5557 | -20.9114
12.8251 | -34.5561 | -20.9198
12.8312 | -34.5565 | -20.9283
12.8373 | -34.557 | -20.9369
12.8434 | -34.5575 | -20.9453
12.8495 | -34.5579 | -20.9537
12.8556 | -34.5583 | -20.9621
12.8617 | -34.5587 | -20.9703
12.8678 | -34.5591 | -20.9788
12.8739 | -34.5596 | -20.987
12.88 | -34.56 | -20.9955
12.8861 | -34.5604 | -21.0041
12.8922 | -34.5608 | -21.0126
12.8983 | -34.5613 | -21.0207
12.9044 | -34.5618 | -21.0291
12.9105 | -34.5623 | -21.0375
12.9166 | -34.5627 | -21.0463
12.9227 | -34.5631 | -21.0553
12.9288 | -34.5636 | -21.0638
12.9349 | -34.5641 | -21.0726
12.941 | -34.5646 | -21.0812
12.9471 | -34.5651 | -21.0896
12.9532 | -34.5656 | -21.0981
12.9593 | -34.566 | -21.1068
12.9654 | -34.5664 | -21.1155
12.9716 | -34.5668 | -21.124
12.9777 | -34.5672 | -21.1321
12.9838 | -34.5677 | -21.1405
12.9899 | -34.5681 | -21.1491
12.996 | -34.5685 | -21.1578
13.0021 | -34.5689 | -21.1663
13.0082 | -34.5693 | -21.1747
13.0143 | -34.5698 | -21.1829
13.0204 | -34.5701 | -21.1911
13.0265 | -34.5706 | -21.1992
13.0326 | -34.571 | -21.208
13.0387 | -34.5715 | -21.2167
13.0448 | -34.5719 | -21.2251
13.0509 | -34.5724 | -21.2336
13.057 | -34.5729 | -21.2417
13.0631 | -34.5734 | -21.2499
13.0692 | -34.5738 | -21.2582
13.0753 | -34.5743 | -21.2666
13.0814 | -34.5747 | -21.2748
13.0875 | -34.5751 | -21.2829
13.0936 | -34.5756 | -21.2912
13.0997 | -34.576 | -21.2992
13.1058 | -34.5764 | -21.307
13.112 | -34.5768 | -21.3154
13.1181 | -34.5773 | -21.3238
13.1242 | -34.5777 | -21.3323
13.1303 | -34.5782 | -21.3409
13.1364 | -34.5786 | -21.3493
13.1425 | -34.579 | -21.3576
13.1486 | -34.5794 | -21.3658
13.1547 | -34.5799 | -21.3739
13.1608 | -34.5803 | -21.3822
13.1669 | -34.5808 | -21.3904
13.173 | -34.5812 | -21.3988
13.1791 | -34.5816 | -21.4072
13.1852 | -34.5821 | -21.4159
13.1913 | -34.5825 | -21.4243
13.1974 | -34.5829 | -21.4326
13.2035 | -34.5833 | -21.4409
13.2096 | -34.5838 | -21.4493
13.2157 | -34.5843 | -21.4576
13.2218 | -34.5847 | -21.466
13.2279 | -34.5852 | -21.4743
13.234 | -34.5857 | -21.4827
13.2401 | -34.5861 | -21.491
13.2462 | -34.5865 | -21.4993
13.2524 | -34.5869 | -21.5076
13.2585 | -34.5874 | -21.5159
13.2646 | -34.5878 | -21.5245
13.2707 | -34.5883 | -21.5329
13.2768 | -34.5887 | -21.5417
13.2829 | -34.5891 | -21.5501
13.289 | -34.5895 | -21.5584
13.2951 | -34.59 | -21.5666
13.3012 | -34.5904 | -21.5751
13.3073 | -34.5908 | -21.5837
13.3134 | -34.5912 | -21.592
13.3195 | -34.5917 | -21.6001
13.3256 | -34.5921 | -21.6086
13.3317 | -34.5926 | -21.6167
13.3378 | -34.593 | -21.6252
13.3439 | -34.5934 | -21.6336
13.35 | -34.5939 | -21.6419
13.3561 | -34.5943 | -21.6504
13.3622 | -34.5947 | -21.6587
13.3683 | -34.5952 | -21.6668
13.3744 | -34.5956 | -21.675
13.3805 | -34.596 | -21.6834
13.3866 | -34.5964 | -21.6915
13.3927 | -34.5968 | -21.6998
13.3989 | -34.5973 | -21.7079
13.405 | -34.5978 | -21.7161
13.4111 | -34.5983 | -21.7244
13.4172 | -34.5987 | -21.7324
13.4233 | -34.5991 | -21.741
13.4294 | -34.5995 | -21.749
13.4355 | -34.6 | -21.7579
13.4416 | -34.6004 | -21.7662
13.4477 | -34.6008 | -21.7746
13.4538 | -34.6012 | -21.7828
13.4599 | -34.6017 | -21.7913
13.466 | -34.6021 | -21.8
13.4721 | -34.6026 | -21.8084
13.4782 | -34.603 | -21.8166
13.4843 | -34.6035 | -21.8249
13.4904 | -34.6039 | -21.8327
13.4965 | -34.6043 | -21.8415
13.5026 | -34.6048 | -21.8497
13.5087 | -34.6052 | -21.8577
13.5148 | -34.6056 | -21.8661
13.5209 | -34.6061 | -21.8743
13.527 | -34.6065 | -21.8829
13.5331 | -34.6069 | -21.8912
13.5393 | -34.6074 | -21.8995
13.5454 | -34.6078 | -21.9079
13.5515 | -34.6082 | -21.9162
13.5576 | -34.6086 | -21.9246
13.5637 | -34.6091 | -21.9328
13.5698 | -34.6095 | -21.9414
13.5759 | -34.6099 | -21.9492
13.582 | -34.6103 | -21.9577
13.5881 | -34.6108 | -21.966
13.5942 | -34.6112 | -21.9743
13.6003 | -34.6117 | -21.9827
13.6064 | -34.6121 | -21.9912
13.6125 | -34.6125 | -21.9999
13.6186 | -34.613 | -22.0081
13.6247 | -34.6134 | -22.0164
13.6308 | -34.6138 | -22.0244
13.6369 | -34.6143 | -22.0327
13.643 | -34.6147 | -22.0408
13.6491 | -34.6152 | -22.0492
13.6552 | -34.6156 | -22.0573
13.6613 | -34.616 | -22.0658
13.6674 | -34.6165 | -22.0742
13.6735 | -34.6169 | -22.0826
13.6796 | -34.6173 | -22.0907
13.6858 | -34.6177 | -22.0991
13.6919 | -34.6182 | -22.1075
13.698 | -34.6186 | -22.1159
13.7041 | -34.619 | -22.1245
13.7102 | -34.6195 | -22.1329
13.7163 | -34.6199 | -22.1412
13.7224 | -34.6204 | -22.1494
13.7285 | -34.6208 | -22.1578
13.7346 | -34.6212 | -22.1661
13.7407 | -34.6217 | -22.1744
13.7468 | -34.6221 | -22.1828
13.7529 | -34.6225 | -22.1909
13.759 | -34.6229 | -22.1994
13.7651 | -34.6234 | -22.2077
13.7712 | -34.6238 | -22.2163
13.7773 | -34.6242 | -22.2247
13.7834 | -34.6247 | -22.2329
13.7895 | -34.6252 | -22.2411
13.7956 | -34.6256 | -22.2495
13.8017 | -34.6261 | -22.2577
13.8078 | -34.6265 | -22.266
13.8139 | -34.627 | -22.2742
13.82 | -34.6274 | -22.2822
13.8262 | -34.6279 | -22.2909
13.8323 | -34.6283 | -22.2994
13.8384 | -34.6288 | -22.3079
13.8445 | -34.6292 | -22.3163
13.8506 | -34.6296 | -22.3247
13.8567 | -34.6301 | -22.3333
13.8628 | -34.6305 | -22.3417
13.8689 | -34.6309 | -22.3501
13.875 | -34.6314 | -22.3584
13.8811 | -34.6318 | -22.3665
13.8872 | -34.6323 | -22.3747
13.8933 | -34.6328 | -22.3829
13.8994 | -34.6332 | -22.3909
13.9055 | -34.6337 | -22.3992
13.9116 | -34.6341 | -22.4072
13.9177 | -34.6346 | -22.4151
13.9238 | -34.6351 | -22.4236
13.9299 | -34.6355 | -22.4319
13.936 | -34.6359 | -22.4401
13.9421 | -34.6363 | -22.4483
13.9482 | -34.6368 | -22.4566
13.9543 | -34.6373 | -22.4646
13.9604 | -34.6377 | -22.4731
13.9665 | -34.6382 | -22.4815
13.9727 | -34.6386 | -22.4896
13.9788 | -34.639 | -22.4981
13.9849 | -34.6394 | -22.5065
13.991 | -34.6398 | -22.5151
13.9971 | -34.6403 | -22.5236
14.0032 | -34.6407 | -22.5322
14.0093 | -34.6411 | -22.5403
14.0154 | -34.6415 | -22.5488
14.0215 | -34.6419 | -22.5568
14.0276 | -34.6424 | -22.5649
14.0337 | -34.6429 | -22.5735
14.0398 | -34.6433 | -22.5815
14.0459 | -34.6437 | -22.59
14.052 | -34.6441 | -22.5983
14.0581 | -34.6445 | -22.6062
14.0642 | -34.645 | -22.6147
14.0703 | -34.6454 | -22.623
14.0764 | -34.6459 | -22.6316
14.0825 | -34.6463 | -22.64
14.0886 | -34.6467 | -22.6478
14.0947 | -34.6471 | -22.6559
14.1008 | -34.6476 | -22.6645
14.1069 | -34.648 | -22.6727
14.1131 | -34.6484 | -22.6813
14.1192 | -34.6489 | -22.6892
14.1253 | -34.6493 | -22.6972
14.1314 | -34.6498 | -22.7054
14.1375 | -34.6502 | -22.7139
14.1436 | -34.6507 | -22.7221
14.1497 | -34.6512 | -22.7302
14.1558 | -34.6516 | -22.7386
14.1619 | -34.652 | -22.7468
14.168 | -34.6525 | -22.7549
14.1741 | -34.653 | -22.7631
14.1802 | -34.6534 | -22.7712
14.1863 | -34.6538 | -22.7795
14.1924 | -34.6543 | -22.7878
14.1985 | -34.6547 | -22.7959
14.2046 | -34.6552 | -22.8039
14.2107 | -34.6556 | -22.8117
14.2168 | -34.6561 | -22.8203
14.2229 | -34.6565 | -22.8286
14.229 | -34.657 | -22.8369
14.2351 | -34.6574 | -22.8449
14.2412 | -34.6578 | -22.853
14.2473 | -34.6583 | -22.8612
14.2534 | -34.6588 | -22.8694
14.2596 | -34.6592 | -22.8776
14.2657 | -34.6597 | -22.8859
14.2718 | -34.6601 | -22.8943
14.2779 | -34.6606 | -22.9023
14.284 | -34.661 | -22.9106
14.2901 | -34.6615 | -22.9186
14.2962 | -34.662 | -22.9266
14.3023 | -34.6624 | -22.9349
14.3084 | -34.6629 | -22.9434
14.3145 | -34.6633 | -22.9516
14.3206 | -34.6637 | -22.9599
14.3267 | -34.6642 | -22.9681
14.3328 | -34.6646 | -22.9762
14.3389 | -34.6651 | -22.9841
14.345 | -34.6655 | -22.9925
14.3511 | -34.6659 | -23.0004
14.3572 | -34.6664 | -23.0089
14.3633 | -34.6668 | -23.017
14.3694 | -34.6672 | -23.0251
14.3755 | -34.6676 | -23.033
14.3816 | -34.6681 | -23.0412
14.3877 | -34.6686 | -23.0492
14.3938 | -34.669 | -23.0576
14.4 | -34.6695 | -23.0655
14.4061 | -34.6699 | -23.0737
14.4122 | -34.6703 | -23.082
14.4183 | -34.6707 | -23.0902
14.4244 | -34.6711 | -23.0982
14.4305 | -34.6716 | -23.1067
14.4366 | -34.672 | -23.115
14.4427 | -34.6725 | -23.1234
14.4488 | -34.6729 | -23.1313
14.4549 | -34.6734 | -23.1394
14.461 | -34.6738 | -23.1477
14.4671 | -34.6743 | -23.156
14.4732 | -34.6747 | -23.1638
14.4793 | -34.6752 | -23.1726
14.4854 | -34.6757 | -23.1804
14.4915 | -34.6761 | -23.1884
14.4976 | -34.6766 | -23.1966
14.5037 | -34.6771 | -23.2049
14.5098 | -34.6775 | -23.2129
14.5159 | -34.678 | -23.2212
14.522 | -34.6784 | -23.2295
14.5281 | -34.6789 | -23.2374
14.5342 | -34.6793 | -23.2456
14.5403 | -34.6798 | -23.2538
14.5465 | -34.6802 | -23.2624
14.5526 | -34.6807 | -23.2709
14.5587 | -34.6811 | -23.279
14.5648 | -34.6816 | -23.287
14.5709 | -34.682 | -23.2953
14.577 | -34.6824 | -23.3034
14.5831 | -34.6829 | -23.3117
14.5892 | -34.6834 | -23.3196
14.5953 | -34.6838 | -23.3274
14.6014 | -34.6843 | -23.3356
14.6075 | -34.6848 | -23.3434
14.6136 | -34.6852 | -23.3514
14.6197 | -34.6856 | -23.3596
14.6258 | -34.6861 | -23.3678
14.6319 | -34.6866 | -23.3758
14.638 | -34.687 | -23.3839
14.6441 | -34.6875 | -23.3916
14.6502 | -34.6879 | -23.3995
14.6563 | -34.6883 | -23.4075
14.6624 | -34.6888 | -23.4155
14.6685 | -34.6893 | -23.4233
14.6746 | -34.6897 | -23.4314
14.6807 | -34.6901 | -23.4395
14.6869 | -34.6906 | -23.4475
14.693 | -34.691 | -23.4554
14.6991 | -34.6915 | -23.4633
14.7052 | -34.692 | -23.4713
14.7113 | -34.6924 | -23.4797
14.7174 | -34.6928 | -23.4877
14.7235 | -34.6933 | -23.4956
14.7296 | -34.6937 | -23.5036
14.7357 | -34.6942 | -23.5115
14.7418 | -34.6947 | -23.5192
14.7479 | -34.6951 | -23.5273
14.754 | -34.6956 | -23.5354
14.7601 | -34.6961 | -23.5436
14.7662 | -34.6965 | -23.5517
14.7723 | -34.697 | -23.5596
14.7784 | -34.6974 | -23.5678
14.7845 | -34.6979 | -23.5759
14.7906 | -34.6983 | -23.5841
14.7967 | -34.6988 | -23.592
14.8028 | -34.6992 | -23.6
14.8089 | -34.6997 | -23.6079
14.815 | -34.7001 | -23.6158
14.8211 | -34.7005 | -23.6238
14.8272 | -34.701 | -23.632
14.8334 | -34.7014 | -23.6399
14.8395 | -34.7019 | -23.6479
14.8456 | -34.7024 | -23.6558
14.8517 | -34.7028 | -23.664
14.8578 | -34.7033 | -23.672
14.8639 | -34.7038 | -23.6802
14.87 | -34.7042 | -23.6882
14.8761 | -34.7047 | -23.696
14.8822 | -34.7051 | -23.7041
14.8883 | -34.7056 | -23.712
14.8944 | -34.706 | -23.72
14.9005 | -34.7065 | -23.7279
14.9066 | -34.707 | -23.7359
14.9127 | -34.7074 | -23.7437
14.9188 | -34.7079 | -23.7519
14.9249 | -34.7083 | -23.76
14.931 | -34.7088 | -23.7679
14.9371 | -34.7092 | -23.7758
14.9432 | -34.7097 | -23.7839
14.9493 | -34.7101 | -23.7919
14.9554 | -34.7106 | -23.7998
14.9615 | -34.711 | -23.8078
14.9676 | -34.7115 | -23.8159
14.9738 | -34.712 | -23.8237
14.9799 | -34.7124 | -23.8321
14.986 | -34.7129 | -23.8397
14.9921 | -34.7134 | -23.8475
14.9982 | -34.7138 | -23.8558
15.0043 | -34.7143 | -23.8637
15.0104 | -34.7147 | -23.8717
15.0165 | -34.7152 | -23.8798
15.0226 | -34.7157 | -23.8874
15.0287 | -34.7161 | -23.8958
15.0348 | -34.7166 | -23.9039
15.0409 | -34.7171 | -23.9122
15.047 | -34.7175 | -23.9203
15.0531 | -34.7181 | -23.9284
15.0592 | -34.7186 | -23.9366
15.0653 | -34.719 | -23.9445
15.0714 | -34.7195 | -23.9521
15.0775 | -34.72 | -23.9598
15.0836 | -34.7205 | -23.9678
15.0897 | -34.721 | -23.9757
15.0958 | -34.7214 | -23.9835
15.1019 | -34.7218 | -23.9915
15.108 | -34.7223 | -23.9996
15.1141 | -34.7227 | -24.0076
15.1203 | -34.7233 | -24.0156
15.1264 | -34.7238 | -24.0237
15.1325 | -34.7243 | -24.0316
15.1386 | -34.7247 | -24.0397
15.1447 | -34.7252 | -24.0477
15.1508 | -34.7257 | -24.0559
15.1569 | -34.7262 | -24.0642
15.163 | -34.7267 | -24.0717
15.1691 | -34.7272 | -24.0796
15.1752 | -34.7277 | -24.0875
15.1813 | -34.7282 | -24.0954
15.1874 | -34.7286 | -24.1034
15.1935 | -34.7292 | -24.111
15.1996 | -34.7297 | -24.1192
15.2057 | -34.7301 | -24.1271
15.2118 | -34.7306 | -24.1349
15.2179 | -34.7311 | -24.1426
15.224 | -34.7315 | -24.1507
15.2301 | -34.732 | -24.1582
15.2362 | -34.7325 | -24.1661
15.2423 | -34.733 | -24.1744
15.2484 | -34.7334 | -24.1822
15.2545 | -34.7339 | -24.19
15.2607 | -34.7344 | -24.198
15.2668 | -34.735 | -24.2056
15.2729 | -34.7354 | -24.2134
15.279 | -34.7359 | -24.2213
15.2851 | -34.7364 | -24.2291
15.2912 | -34.7369 | -24.2369
15.2973 | -34.7373 | -24.2447
15.3034 | -34.7378 | -24.2524
15.3095 | -34.7383 | -24.26
15.3156 | -34.7387 | -24.268
15.3217 | -34.7392 | -24.2755
15.3278 | -34.7397 | -24.2836
15.3339 | -34.7403 | -24.2912
15.34 | -34.7407 | -24.2987
15.3461 | -34.7412 | -24.3066
15.3522 | -34.7416 | -24.3142
15.3583 | -34.7421 | -24.322
15.3644 | -34.7426 | -24.33
15.3705 | -34.7431 | -24.3376
15.3766 | -34.7436 | -24.3458
15.3827 | -34.744 | -24.3534
15.3888 | -34.7445 | -24.361
15.3949 | -34.745 | -24.3687
15.401 | -34.7455 | -24.3766
15.4072 | -34.746 | -24.3843
15.4133 | -34.7465 | -24.3921
15.4194 | -34.747 | -24.4
15.4255 | -34.7474 | -24.4079
15.4316 | -34.748 | -24.4158
15.4377 | -34.7484 | -24.4236
15.4438 | -34.7489 | -24.4311
15.4499 | -34.7494 | -24.4389
15.456 | -34.7498 | -24.4466
15.4621 | -34.7503 | -24.4545
15.4682 | -34.7508 | -24.4626
15.4743 | -34.7514 | -24.4702
15.4804 | -34.7518 | -24.4781
15.4865 | -34.7523 | -24.4858
15.4926 | -34.7527 | -24.4938
15.4987 | -34.7532 | -24.5017
15.5048 | -34.7537 | -24.5095
15.5109 | -34.7542 | -24.5169
15.517 | -34.7547 | -24.5247
15.5231 | -34.7552 | -24.5326
15.5292 | -34.7557 | -24.5402
15.5353 | -34.7562 | -24.548
15.5414 | -34.7567 | -24.556
15.5476 | -34.7572 | -24.5639
15.5537 | -34.7577 | -24.5716
15.5598 | -34.7582 | -24.5791
15.5659 | -34.7587 | -24.5871
15.572 | -34.7592 | -24.5945
15.5781 | -34.7597 | -24.6023
15.5842 | -34.7602 | -24.6098
15.5903 | -34.7607 | -24.6172
15.5964 | -34.7612 | -24.6247
15.6025 | -34.7617 | -24.6323
15.6086 | -34.7622 | -24.6402
15.6147 | -34.7626 | -24.6478
15.6208 | -34.7632 | -24.6553
15.6269 | -34.7637 | -24.6629
15.633 | -34.7641 | -24.6706
15.6391 | -34.7646 | -24.6786
15.6452 | -34.7651 | -24.6862
15.6513 | -34.7656 | -24.6942
15.6574 | -34.7661 | -24.7017
15.6635 | -34.7666 | -24.7093
15.6696 | -34.7671 | -24.7169
15.6757 | -34.7676 | -24.7249
15.6818 | -34.7681 | -24.733
15.688 | -34.7686 | -24.7407
15.6941 | -34.7691 | -24.7483
15.7002 | -34.7696 | -24.756
15.7063 | -34.7701 | -24.7638
15.7124 | -34.7706 | -24.7716
15.7185 | -34.7711 | -24.7792
15.7246 | -34.7716 | -24.7867
15.7307 | -34.7721 | -24.7945
15.7368 | -34.7726 | -24.8026
15.7429 | -34.7731 | -24.8101
15.749 | -34.7735 | -24.8179
15.7551 | -34.774 | -24.8256
15.7612 | -34.7745 | -24.8327
15.7673 | -34.775 | -24.8402
15.7734 | -34.7754 | -24.8478
15.7795 | -34.7759 | -24.8559
15.7856 | -34.7765 | -24.8633
15.7917 | -34.777 | -24.8708
15.7978 | -34.7775 | -24.8785
15.8039 | -34.7779 | -24.886
15.81 | -34.7784 | -24.8936
15.8161 | -34.7789 | -24.9013
15.8222 | -34.7794 | -24.909
15.8283 | -34.7799 | -24.9165
15.8345 | -34.7804 | -24.9242
15.8406 | -34.781 | -24.9318
15.8467 | -34.7815 | -24.9395
15.8528 | -34.782 | -24.9471
15.8589 | -34.7824 | -24.9546
15.865 | -34.7829 | -24.9621
15.8711 | -34.7834 | -24.97
15.8772 | -34.784 | -24.9777
15.8833 | -34.7845 | -24.9857
15.8894 | -34.785 | -24.9933
15.8955 | -34.7855 | -25.0011
15.9016 | -34.786 | -25.0086
15.9077 | -34.7865 | -25.0162
15.9138 | -34.787 | -25.0238
15.9199 | -34.7875 | -25.0316
15.926 | -34.788 | -25.0395
15.9321 | -34.7885 | -25.0471
15.9382 | -34.7891 | -25.0547
15.9443 | -34.7896 | -25.0626
15.9504 | -34.7901 | -25.0702
15.9565 | -34.7907 | -25.0779
15.9626 | -34.7912 | -25.0855
15.9687 | -34.7917 | -25.093
15.9749 | -34.7922 | -25.1007
15.981 | -34.7927 | -25.1083
15.9871 | -34.7932 | -25.116
15.9932 | -34.7937 | -25.1238
15.9993 | -34.7942 | -25.1314
16.0054 | -34.7947 | -25.1387
16.0115 | -34.7952 | -25.1462
16.0176 | -34.7957 | -25.154
16.0237 | -34.7962 | -25.1621
16.0298 | -34.7967 | -25.1691
16.0359 | -34.7972 | -25.1766
16.042 | -34.7978 | -25.1842
16.0481 | -34.7983 | -25.1918
16.0542 | -34.7988 | -25.1998
16.0603 | -34.7993 | -25.2073
16.0664 | -34.7998 | -25.215
16.0725 | -34.8004 | -25.2223
16.0786 | -34.8009 | -25.2302
16.0847 | -34.8014 | -25.2376
16.0908 | -34.8019 | -25.2451
16.0969 | -34.8024 | -25.2525
16.103 | -34.8028 | -25.2604
16.1091 | -34.8034 | -25.2678
16.1152 | -34.8038 | -25.2755
16.1214 | -34.8044 | -25.2832
16.1275 | -34.8049 | -25.2906
16.1336 | -34.8054 | -25.2982
16.1397 | -34.8059 | -25.3057
16.1458 | -34.8065 | -25.3132
16.1519 | -34.8069 | -25.321
16.158 | -34.8075 | -25.3287
16.1641 | -34.808 | -25.3364
16.1702 | -34.8085 | -25.3441
16.1763 | -34.809 | -25.3518
16.1824 | -34.8095 | -25.3591
16.1885 | -34.81 | -25.3667
16.1946 | -34.8105 | -25.3742
16.2007 | -34.811 | -25.3822
16.2068 | -34.8115 | -25.3897
16.2129 | -34.812 | -25.3975
16.219 | -34.8125 | -25.405
16.2251 | -34.813 | -25.4128
16.2312 | -34.8134 | -25.4205
16.2373 | -34.8139 | -25.4281
16.2434 | -34.8144 | -25.4356
16.2495 | -34.815 | -25.443
16.2556 | -34.8155 | -25.4506
16.2618 | -34.816 | -25.4579
16.2679 | -34.8166 | -25.4655
16.274 | -34.8171 | -25.4733
16.2801 | -34.8176 | -25.4806
16.2862 | -34.8181 | -25.4881
16.2923 | -34.8186 | -25.4959
16.2984 | -34.8191 | -25.5036
16.3045 | -34.8196 | -25.5111
16.3106 | -34.8202 | -25.5183
16.3167 | -34.8206 | -25.5261
16.3228 | -34.8212 | -25.5338
16.3289 | -34.8217 | -25.5413
16.335 | -34.8222 | -25.5491
16.3411 | -34.8226 | -25.5563
16.3472 | -34.8232 | -25.5641
16.3533 | -34.8237 | -25.5718
16.3594 | -34.8241 | -25.5791
16.3655 | -34.8246 | -25.5868
16.3716 | -34.8251 | -25.5942
16.3777 | -34.8257 | -25.6017
16.3838 | -34.8261 | -25.6092
16.3899 | -34.8267 | -25.6168
16.396 | -34.8272 | -25.6245
16.4021 | -34.8277 | -25.6319
16.4083 | -34.8282 | -25.6394
16.4144 | -34.8287 | -25.6467
16.4205 | -34.8293 | -25.6544
16.4266 | -34.8297 | -25.6619
16.4327 | -34.8303 | -25.6692
16.4388 | -34.8307 | -25.6768
16.4449 | -34.8313 | -25.6843
16.451 | -34.8318 | -25.692
16.4571 | -34.8322 | -25.6997
16.4632 | -34.8327 | -25.7071
16.4693 | -34.8332 | -25.7148
16.4754 | -34.8337 | -25.7224
16.4815 | -34.8342 | -25.7301
16.4876 | -34.8347 | -25.7376
16.4937 | -34.8353 | -25.7451
16.4998 | -34.8358 | -25.7526
16.5059 | -34.8363 | -25.7601
16.512 | -34.8368 | -25.768
16.5181 | -34.8374 | -25.7756
16.5242 | -34.8379 | -25.7833
16.5303 | -34.8383 | -25.7906
16.5364 | -34.8388 | -25.798
16.5425 | -34.8393 | -25.8056
16.5487 | -34.8398 | -25.813
16.5548 | -34.8403 | -25.8205
16.5609 | -34.8408 | -25.8282
16.567 | -34.8413 | -25.8357
16.5731 | -34.8418 | -25.8433
16.5792 | -34.8423 | -25.8509
16.5853 | -34.8429 | -25.8583
16.5914 | -34.8434 | -25.866
16.5975 | -34.8439 | -25.8736
16.6036 | -34.8444 | -25.8811
16.6097 | -34.8449 | -25.8885
16.6158 | -34.8454 | -25.8962
16.6219 | -34.8459 | -25.9036
16.628 | -34.8464 | -25.9111
16.6341 | -34.8469 | -25.9185
16.6402 | -34.8474 | -25.9256
16.6463 | -34.8479 | -25.9331
16.6524 | -34.8484 | -25.9405
16.6585 | -34.8489 | -25.9482
16.6646 | -34.8494 | -25.9558
16.6707 | -34.8499 | -25.9632
16.6768 | -34.8504 | -25.9706
16.6829 | -34.8509 | -25.978
16.689 | -34.8514 | -25.9856
16.6952 | -34.8519 | -25.9932
16.7013 | -34.8524 | -26.0008
16.7074 | -34.8529 | -26.0083
16.7135 | -34.8534 | -26.0156
16.7196 | -34.8539 | -26.0232
16.7257 | -34.8544 | -26.0306
16.7318 | -34.8549 | -26.0381
16.7379 | -34.8554 | -26.0457
16.744 | -34.856 | -26.053
16.7501 | -34.8565 | -26.0605
16.7562 | -34.857 | -26.068
16.7623 | -34.8575 | -26.0754
16.7684 | -34.858 | -26.0831
16.7745 | -34.8585 | -26.0903
16.7806 | -34.859 | -26.0979
16.7867 | -34.8595 | -26.1053
16.7928 | -34.86 | -26.1128
16.7989 | -34.8605 | -26.1201
16.805 | -34.8609 | -26.1277
16.8111 | -34.8615 | -26.1351
16.8172 | -34.8619 | -26.1429
16.8233 | -34.8625 | -26.1506
16.8294 | -34.863 | -26.1582
16.8356 | -34.8635 | -26.1659
16.8417 | -34.864 | -26.1733
16.8478 | -34.8645 | -26.1807
16.8539 | -34.865 | -26.1881
16.86 | -34.8655 | -26.1956
16.8661 | -34.866 | -26.2031
16.8722 | -34.8665 | -26.2106
16.8783 | -34.867 | -26.2178
16.8844 | -34.8675 | -26.2252
16.8905 | -34.8681 | -26.2329
16.8966 | -34.8685 | -26.2404
16.9027 | -34.869 | -26.2477
16.9088 | -34.8695 | -26.2556
16.9149 | -34.8701 | -26.2634
16.921 | -34.8706 | -26.2707
16.9271 | -34.8711 | -26.2785
16.9332 | -34.8716 | -26.286
16.9393 | -34.8721 | -26.2938
16.9454 | -34.8726 | -26.3011
16.9515 | -34.873 | -26.3088
16.9576 | -34.8735 | -26.3162
16.9637 | -34.8741 | -26.3237
16.9698 | -34.8746 | -26.331
16.9759 | -34.8751 | -26.3382
16.9821 | -34.8756 | -26.3459
16.9882 | -34.8761 | -26.3534
16.9943 | -34.8766 | -26.361
17.0004 | -34.8771 | -26.3686
17.0065 | -34.8776 | -26.3761
17.0126 | -34.8781 | -26.3835
17.0187 | -34.8786 | -26.3911
17.0248 | -34.8791 | -26.3982
17.0309 | -34.8796 | -26.4057
17.037 | -34.8801 | -26.4133
17.0431 | -34.8806 | -26.4204
17.0492 | -34.8811 | -26.428
17.0553 | -34.8816 | -26.4354
17.0614 | -34.8821 | -26.4429
17.0675 | -34.8826 | -26.4506
17.0736 | -34.8831 | -26.4581
17.0797 | -34.8836 | -26.4651
17.0858 | -34.8841 | -26.4728
17.0919 | -34.8845 | -26.4805
17.098 | -34.8851 | -26.4878
17.1041 | -34.8856 | -26.495
17.1102 | -34.8861 | -26.5027
17.1163 | -34.8866 | -26.5103
17.1225 | -34.8871 | -26.5176
17.1286 | -34.8876 | -26.5248
17.1347 | -34.8881 | -26.5322
17.1408 | -34.8886 | -26.5396
17.1469 | -34.8891 | -26.547
17.153 | -34.8897 | -26.5544
17.1591 | -34.8901 | -26.5617
17.1652 | -34.8906 | -26.5689
17.1713 | -34.8911 | -26.5763
17.1774 | -34.8916 | -26.5838
17.1835 | -34.8921 | -26.5914
17.1896 | -34.8926 | -26.5987
17.1957 | -34.8931 | -26.6062
17.2018 | -34.8936 | -26.6137
17.2079 | -34.8941 | -26.6211
17.214 | -34.8946 | -26.6284
17.2201 | -34.8951 | -26.6359
17.2262 | -34.8955 | -26.6434
17.2323 | -34.896 | -26.6508
17.2384 | -34.8966 | -26.6582
17.2445 | -34.8971 | -26.6656
17.2506 | -34.8975 | -26.673
17.2567 | -34.898 | -26.6805
17.2628 | -34.8985 | -26.6881
17.269 | -34.899 | -26.6955
17.2751 | -34.8995 | -26.7029
17.2812 | -34.9 | -26.7103
17.2873 | -34.9005 | -26.7179
17.2934 | -34.9009 | -26.7252
17.2995 | -34.9014 | -26.7325
17.3056 | -34.9019 | -26.7398
17.3117 | -34.9024 | -26.7472
17.3178 | -34.9029 | -26.7545
17.3239 | -34.9033 | -26.7618
17.33 | -34.9038 | -26.7691
17.3361 | -34.9043 | -26.7764
17.3422 | -34.9048 | -26.7838
17.3483 | -34.9053 | -26.7913
17.3544 | -34.9058 | -26.7986
17.3605 | -34.9063 | -26.806
17.3666 | -34.9068 | -26.8137
17.3727 | -34.9073 | -26.8211
17.3788 | -34.9078 | -26.8283
17.3849 | -34.9083 | -26.8357
17.391 | -34.9088 | -26.8429
17.3971 | -34.9093 | -26.8504
17.4032 | -34.9098 | -26.8582
17.4094 | -34.9103 | -26.8656
17.4155 | -34.9107 | -26.8728
17.4216 | -34.9112 | -26.8802
17.4277 | -34.9117 | -26.8878
17.4338 | -34.9122 | -26.8952
17.4399 | -34.9127 | -26.9027
17.446 | -34.9132 | -26.91
17.4521 | -34.9137 | -26.9172
17.4582 | -34.9142 | -26.9246
17.4643 | -34.9147 | -26.932
17.4704 | -34.9151 | -26.9392
17.4765 | -34.9156 | -26.9466
17.4826 | -34.9161 | -26.954
17.4887 | -34.9166 | -26.9616
17.4948 | -34.9171 | -26.9689
17.5009 | -34.9176 | -26.9761
17.507 | -34.9181 | -26.9836
17.5131 | -34.9186 | -26.9909
17.5192 | -34.919 | -26.9983
17.5253 | -34.9195 | -27.0058
17.5314 | -34.9199 | -27.0131
17.5375 | -34.9204 | -27.0206
17.5436 | -34.9209 | -27.0281
17.5497 | -34.9214 | -27.0357
17.5559 | -34.9218 | -27.0427
17.562 | -34.9223 | -27.05
17.5681 | -34.9228 | -27.0573
17.5742 | -34.9233 | -27.0647
17.5803 | -34.9238 | -27.0719
17.5864 | -34.9244 | -27.079
17.5925 | -34.9249 | -27.0867
17.5986 | -34.9254 | -27.0937
17.6047 | -34.9259 | -27.1014
17.6108 | -34.9264 | -27.1085
17.6169 | -34.9269 | -27.1157
17.623 | -34.9273 | -27.123
17.6291 | -34.9278 | -27.1302
17.6352 | -34.9283 | -27.1374
17.6413 | -34.9288 | -27.1449
17.6474 | -34.9293 | -27.1524
17.6535 | -34.9298 | -27.1598
17.6596 | -34.9303 | -27.1671
17.6657 | -34.9307 | -27.1743
17.6718 | -34.9312 | -27.1818
17.6779 | -34.9317 | -27.189
17.684 | -34.9322 | -27.1964
17.6901 | -34.9327 | -27.2036
17.6963 | -34.9332 | -27.2108
17.7024 | -34.9337 | -27.2181
17.7085 | -34.9341 | -27.2252
17.7146 | -34.9346 | -27.2324
17.7207 | -34.9351 | -27.2397
17.7268 | -34.9356 | -27.2468
17.7329 | -34.9361 | -27.2545
17.739 | -34.9366 | -27.262
17.7451 | -34.9371 | -27.2692
17.7512 | -34.9376 | -27.2766
17.7573 | -34.938 | -27.284
17.7634 | -34.9385 | -27.291
17.7695 | -34.939 | -27.2984
17.7756 | -34.9395 | -27.3057
17.7817 | -34.94 | -27.3128
17.7878 | -34.9405 | -27.3201
17.7939 | -34.941 | -27.3274
17.8 | -34.9415 | -27.3347
17.8061 | -34.9419 | -27.3421
17.8122 | -34.9424 | -27.3494
17.8183 | -34.9429 | -27.3568
17.8244 | -34.9434 | -27.3641
17.8305 | -34.9439 | -27.3714
17.8366 | -34.9444 | -27.3783
17.8428 | -34.9449 | -27.3859
17.8489 | -34.9454 | -27.393
17.855 | -34.9459 | -27.4001
17.8611 | -34.9464 | -27.4073
17.8672 | -34.9469 | -27.4144
17.8733 | -34.9474 | -27.422
17.8794 | -34.9479 | -27.4294
17.8855 | -34.9484 | -27.4366
17.8916 | -34.9488 | -27.4437
17.8977 | -34.9493 | -27.451
17.9038 | -34.9498 | -27.458
17.9099 | -34.9503 | -27.4653
17.916 | -34.9508 | -27.4725
17.9221 | -34.9512 | -27.4797
17.9282 | -34.9517 | -27.487
17.9343 | -34.9522 | -27.4944
17.9404 | -34.9527 | -27.5015
17.9465 | -34.9532 | -27.5088
17.9526 | -34.9537 | -27.5162
17.9587 | -34.9542 | -27.5235
17.9648 | -34.9546 | -27.5307
17.9709 | -34.9551 | -27.538
17.977 | -34.9556 | -27.545
17.9832 | -34.9561 | -27.5524
17.9893 | -34.9565 | -27.5597
17.9954 | -34.957 | -27.5672
18.0015 | -34.9575 | -27.5745
18.0076 | -34.958 | -27.5814
18.0137 | -34.9585 | -27.589
18.0198 | -34.9589 | -27.5963
18.0259 | -34.9594 | -27.6037
18.032 | -34.9598 | -27.6107
18.0381 | -34.9603 | -27.6178
18.0442 | -34.9608 | -27.6251
18.0503 | -34.9613 | -27.6325
18.0564 | -34.9617 | -27.6397
18.0625 | -34.9622 | -27.6471
18.0686 | -34.9627 | -27.6543
18.0747 | -34.9632 | -27.6616
18.0808 | -34.9636 | -27.6687
18.0869 | -34.9641 | -27.676
18.093 | -34.9646 | -27.6832
18.0991 | -34.9651 | -27.6907
18.1052 | -34.9656 | -27.698
18.1113 | -34.9661 | -27.7051
18.1174 | -34.9665 | -27.7123
18.1236 | -34.9671 | -27.7196
18.1297 | -34.9675 | -27.7268
18.1358 | -34.9681 | -27.7341
18.1419 | -34.9686 | -27.7414
18.148 | -34.9691 | -27.7487
18.1541 | -34.9696 | -27.7557
18.1602 | -34.9701 | -27.7631
18.1663 | -34.9707 | -27.7703
18.1724 | -34.9711 | -27.7778
18.1785 | -34.9716 | -27.7851
18.1846 | -34.972 | -27.7921
18.1907 | -34.9725 | -27.7997
18.1968 | -34.9729 | -27.8071
18.2029 | -34.9734 | -27.814
18.209 | -34.9739 | -27.8215
18.2151 | -34.9744 | -27.829
18.2212 | -34.9749 | -27.8359
18.2273 | -34.9755 | -27.843
18.2334 | -34.976 | -27.8503
18.2395 | -34.9764 | -27.8573
18.2456 | -34.9769 | -27.8644
18.2517 | -34.9773 | -27.8717
18.2578 | -34.9778 | -27.8787
18.2639 | -34.9783 | -27.8861
18.2701 | -34.9788 | -27.8932
18.2762 | -34.9793 | -27.9007
18.2823 | -34.9798 | -27.908
18.2884 | -34.9803 | -27.9153
18.2945 | -34.9808 | -27.9222
18.3006 | -34.9813 | -27.9295
18.3067 | -34.9818 | -27.9367
18.3128 | -34.9823 | -27.9438
18.3189 | -34.9828 | -27.9512
18.325 | -34.9833 | -27.958
18.3311 | -34.9837 | -27.9654
18.3372 | -34.9842 | -27.9723
18.3433 | -34.9847 | -27.9793
18.3494 | -34.9851 | -27.9865
18.3555 | -34.9856 | -27.9935
18.3616 | -34.9861 | -28.0005
18.3677 | -34.9867 | -28.0075
18.3738 | -34.9871 | -28.0146
18.3799 | -34.9876 | -28.0216
18.386 | -34.988 | -28.0287
18.3921 | -34.9885 | -28.0358
18.3982 | -34.989 | -28.0429
18.4043 | -34.9894 | -28.05
18.4105 | -34.9899 | -28.0572
18.4166 | -34.9904 | -28.0643
18.4227 | -34.9909 | -28.0715
18.4288 | -34.9914 | -28.0788
18.4349 | -34.9919 | -28.0859
18.441 | -34.9924 | -28.093
18.4471 | -34.9929 | -28.1002
18.4532 | -34.9933 | -28.1074
18.4593 | -34.9938 | -28.1144
18.4654 | -34.9943 | -28.1213
18.4715 | -34.9948 | -28.1284
18.4776 | -34.9953 | -28.1358
18.4837 | -34.9958 | -28.143
18.4898 | -34.9963 | -28.1504
18.4959 | -34.9968 | -28.1576
18.502 | -34.9973 | -28.1649
18.5081 | -34.9978 | -28.1723
18.5142 | -34.9983 | -28.1797
18.5203 | -34.9988 | -28.1865
18.5264 | -34.9992 | -28.1934
18.5325 | -34.9998 | -28.2007
18.5386 | -35.0003 | -28.2081
18.5447 | -35.0008 | -28.2151
18.5508 | -35.0013 | -28.2221
18.557 | -35.0018 | -28.2294
18.5631 | -35.0022 | -28.2367
18.5692 | -35.0027 | -28.2437
18.5753 | -35.0032 | -28.2506
18.5814 | -35.0037 | -28.2579
18.5875 | -35.0042 | -28.2653
18.5936 | -35.0047 | -28.2728
18.5997 | -35.0052 | -28.28
18.6058 | -35.0056 | -28.2871
18.6119 | -35.0061 | -28.2944
18.618 | -35.0066 | -28.3016
18.6241 | -35.0072 | -28.3089
18.6302 | -35.0077 | -28.3163
18.6363 | -35.0081 | -28.3232
18.6424 | -35.0086 | -28.3301
18.6485 | -35.0091 | -28.3372
18.6546 | -35.0096 | -28.3441
18.6607 | -35.01 | -28.3512
18.6668 | -35.0106 | -28.3582
18.6729 | -35.011 | -28.3655
18.679 | -35.0116 | -28.3724
18.6851 | -35.012 | -28.3794
18.6912 | -35.0125 | -28.3865
18.6974 | -35.013 | -28.3936
18.7035 | -35.0135 | -28.4006
18.7096 | -35.014 | -28.4079
18.7157 | -35.0145 | -28.4149
18.7218 | -35.015 | -28.4219
18.7279 | -35.0155 | -28.4289
18.734 | -35.016 | -28.4359
18.7401 | -35.0165 | -28.4427
18.7462 | -35.017 | -28.4493
18.7523 | -35.0176 | -28.4563
18.7584 | -35.0181 | -28.4636
18.7645 | -35.0186 | -28.4706
18.7706 | -35.0191 | -28.4778
18.7767 | -35.0196 | -28.4847
18.7828 | -35.0201 | -28.4915
18.7889 | -35.0205 | -28.4985
18.795 | -35.0211 | -28.5054
18.8011 | -35.0215 | -28.5127
18.8072 | -35.022 | -28.5199
18.8133 | -35.0225 | -28.5271
18.8194 | -35.023 | -28.534
18.8255 | -35.0235 | -28.541
18.8316 | -35.024 | -28.5482
18.8377 | -35.0245 | -28.555
18.8439 | -35.025 | -28.562
18.85 | -35.0254 | -28.569
18.8561 | -35.0259 | -28.5761
18.8622 | -35.0264 | -28.5833
18.8683 | -35.0269 | -28.5901
18.8744 | -35.0274 | -28.5971
18.8805 | -35.0279 | -28.6042
18.8866 | -35.0284 | -28.6114
18.8927 | -35.0289 | -28.6183
18.8988 | -35.0295 | -28.6254
18.9049 | -35.03 | -28.6324
18.911 | -35.0305 | -28.6394
18.9171 | -35.031 | -28.6464
18.9232 | -35.0315 | -28.6533
18.9293 | -35.032 | -28.6603
18.9354 | -35.0325 | -28.6671
18.9415 | -35.033 | -28.6741
18.9476 | -35.0335 | -28.6812
18.9537 | -35.0339 | -28.6884
18.9598 | -35.0345 | -28.6948
18.9659 | -35.035 | -28.7018
18.972 | -35.0355 | -28.7088
18.9781 | -35.0359 | -28.7157
18.9843 | -35.0364 | -28.7228
18.9904 | -35.0369 | -28.7299
18.9965 | -35.0374 | -28.7369
19.0026 | -35.0379 | -28.7441
19.0087 | -35.0384 | -28.7512
19.0148 | -35.0389 | -28.758
19.0209 | -35.0394 | -28.7653
19.027 | -35.0399 | -28.7725
19.0331 | -35.0404 | -28.7796
19.0392 | -35.0409 | -28.7866
19.0453 | -35.0414 | -28.7937
19.0514 | -35.0419 | -28.8006
19.0575 | -35.0424 | -28.8076
19.0636 | -35.0429 | -28.8142
19.0697 | -35.0434 | -28.8214
19.0758 | -35.0439 | -28.8284
19.0819 | -35.0444 | -28.8356
19.088 | -35.0449 | -28.8423
19.0941 | -35.0454 | -28.8491
19.1002 | -35.0459 | -28.8558
19.1063 | -35.0464 | -28.8628
19.1124 | -35.047 | -28.8697
19.1185 | -35.0475 | -28.8768
19.1246 | -35.0479 | -28.8837
19.1308 | -35.0484 | -28.8903
19.1369 | -35.049 | -28.8971
19.143 | -35.0494 | -28.9041
19.1491 | -35.05 | -28.911
19.1552 | -35.0504 | -28.9182
19.1613 | -35.0509 | -28.9252
19.1674 | -35.0514 | -28.9323
19.1735 | -35.0519 | -28.9394
19.1796 | -35.0524 | -28.9462
19.1857 | -35.0529 | -28.9533
19.1918 | -35.0534 | -28.9603
19.1979 | -35.0539 | -28.9674
19.204 | -35.0544 | -28.9741
19.2101 | -35.0549 | -28.9811
19.2162 | -35.0554 | -28.9883
19.2223 | -35.0559 | -28.9949
19.2284 | -35.0564 | -29.002
19.2345 | -35.0569 | -29.009
19.2406 | -35.0574 | -29.0159
19.2467 | -35.0579 | -29.023
19.2528 | -35.0584 | -29.03
19.2589 | -35.0588 | -29.0369
19.265 | -35.0594 | -29.0438
19.2712 | -35.0598 | -29.0507
19.2773 | -35.0603 | -29.0576
19.2834 | -35.0608 | -29.0644
19.2895 | -35.0613 | -29.0715
19.2956 | -35.0619 | -29.0787
19.3017 | -35.0623 | -29.0855
19.3078 | -35.0628 | -29.0924
19.3139 | -35.0633 | -29.0992
19.32 | -35.0639 | -29.106
19.3261 | -35.0643 | -29.1128
19.3322 | -35.0649 | -29.1198
19.3383 | -35.0654 | -29.1266
19.3444 | -35.0659 | -29.1333
19.3505 | -35.0664 | -29.1402
19.3566 | -35.0668 | -29.1471
19.3627 | -35.0674 | -29.1541
19.3688 | -35.0679 | -29.1612
19.3749 | -35.0684 | -29.1682
19.381 | -35.0689 | -29.1747
19.3871 | -35.0694 | -29.1816
19.3932 | -35.0699 | -29.1887
19.3993 | -35.0703 | -29.1956
19.4054 | -35.0709 | -29.2028
19.4115 | -35.0714 | -29.2099
19.4177 | -35.0719 | -29.2167
19.4238 | -35.0724 | -29.2235
19.4299 | -35.0729 | -29.2303
19.436 | -35.0734 | -29.2375
19.4421 | -35.0739 | -29.244
19.4482 | -35.0745 | -29.2509
19.4543 | -35.075 | -29.2573
19.4604 | -35.0754 | -29.264
19.4665 | -35.076 | -29.271
19.4726 | -35.0765 | -29.2776
19.4787 | -35.077 | -29.2846
19.4848 | -35.0774 | -29.2914
19.4909 | -35.0779 | -29.2983
19.497 | -35.0784 | -29.3051
19.5031 | -35.0789 | -29.3119
19.5092 | -35.0794 | -29.3188
19.5153 | -35.0799 | -29.3258
19.5214 | -35.0804 | -29.3325
19.5275 | -35.0809 | -29.3392
19.5336 | -35.0814 | -29.346
19.5397 | -35.0819 | -29.3529
19.5458 | -35.0824 | -29.3597
19.5519 | -35.0829 | -29.3664
19.5581 | -35.0834 | -29.373
19.5642 | -35.0839 | -29.3798
19.5703 | -35.0844 | -29.3865
19.5764 | -35.0849 | -29.3935
19.5825 | -35.0854 | -29.4004
19.5886 | -35.0859 | -29.4072
19.5947 | -35.0864 | -29.4139
19.6008 | -35.0869 | -29.4208
19.6069 | -35.0874 | -29.4277
19.613 | -35.0879 | -29.4347
19.6191 | -35.0884 | -29.4416
19.6252 | -35.0889 | -29.4485
19.6313 | -35.0894 | -29.4553
19.6374 | -35.0899 | -29.4625
19.6435 | -35.0904 | -29.4692
19.6496 | -35.0909 | -29.4761
19.6557 | -35.0914 | -29.4829
19.6618 | -35.0919 | -29.4895
19.6679 | -35.0924 | -29.4965
19.674 | -35.0929 | -29.5031
19.6801 | -35.0933 | -29.5098
19.6862 | -35.0938 | -29.5166
19.6923 | -35.0944 | -29.5233
19.6984 | -35.0949 | -29.53
19.7046 | -35.0954 | -29.5369
19.7107 | -35.0959 | -29.5435
19.7168 | -35.0964 | -29.5506
19.7229 | -35.0968 | -29.5575
19.729 | -35.0973 | -29.5643
19.7351 | -35.0978 | -29.5708
19.7412 | -35.0983 | -29.5776
19.7473 | -35.0988 | -29.5845
19.7534 | -35.0993 | -29.5915
19.7595 | -35.0998 | -29.5983
19.7656 | -35.1003 | -29.6055
19.7717 | -35.1008 | -29.6119
19.7778 | -35.1013 | -29.6189
19.7839 | -35.1018 | -29.6254
19.79 | -35.1023 | -29.632
19.7961 | -35.1028 | -29.6392
19.8022 | -35.1032 | -29.6458
19.8083 | -35.1037 | -29.6526
19.8144 | -35.1042 | -29.6594
19.8205 | -35.1047 | -29.6662
19.8266 | -35.1052 | -29.6728
19.8327 | -35.1057 | -29.6797
19.8388 | -35.1062 | -29.6863
19.845 | -35.1067 | -29.6929
19.8511 | -35.1072 | -29.6999
19.8572 | -35.1077 | -29.7068
19.8633 | -35.1081 | -29.7135
19.8694 | -35.1087 | -29.7205
19.8755 | -35.1092 | -29.7271
19.8816 | -35.1097 | -29.7339
19.8877 | -35.1102 | -29.7407
19.8938 | -35.1107 | -29.7474
19.8999 | -35.1111 | -29.754
19.906 | -35.1116 | -29.7608
19.9121 | -35.1121 | -29.7673
19.9182 | -35.1127 | -29.7738
19.9243 | -35.1132 | -29.7808
19.9304 | -35.1137 | -29.7878
19.9365 | -35.1141 | -29.795
19.9426 | -35.1146 | -29.8016
19.9487 | -35.1152 | -29.808
19.9548 | -35.1156 | -29.815
19.9609 | -35.1162 | -29.8216
19.967 | -35.1167 | -29.8282
19.9731 | -35.1171 | -29.8349
19.9792 | -35.1176 | -29.8412
19.9853 | -35.1181 | -29.8479
19.9915 | -35.1187 | -29.8546
19.9976 | -35.1192 | -29.8613
20.0037 | -35.1197 | -29.8681
20.0098 | -35.1202 | -29.8747
20.0159 | -35.1207 | -29.8813
20.022 | -35.1212 | -29.8883
20.0281 | -35.1218 | -29.8952
20.0342 | -35.1222 | -29.9017
20.0403 | -35.1227 | -29.9081
20.0464 | -35.1232 | -29.9152
20.0525 | -35.1237 | -29.9215
20.0586 | -35.1243 | -29.9279
20.0647 | -35.1248 | -29.9349
20.0708 | -35.1253 | -29.9419
20.0769 | -35.1257 | -29.9485
20.083 | -35.1262 | -29.9549
20.0891 | -35.1267 | -29.9618
20.0952 | -35.1272 | -29.9683
20.1013 | -35.1276 | -29.975
20.1074 | -35.1281 | -29.9819
20.1135 | -35.1287 | -29.9884
20.1196 | -35.1292 | -29.9951
20.1257 | -35.1297 | -30.0019
20.1319 | -35.1302 | -30.0087
20.138 | -35.1307 | -30.0153
20.1441 | -35.1312 | -30.0218
20.1502 | -35.1317 | -30.0287
20.1563 | -35.1322 | -30.0351
20.1624 | -35.1327 | -30.0418
20.1685 | -35.1332 | -30.0484
20.1746 | -35.1337 | -30.0551
20.1807 | -35.1342 | -30.062
20.1868 | -35.1347 | -30.0689
20.1929 | -35.1352 | -30.0752
20.199 | -35.1357 | -30.0822
20.2051 | -35.1362 | -30.0886
20.2112 | -35.1367 | -30.0953
20.2173 | -35.1372 | -30.1019
20.2234 | -35.1377 | -30.1085
20.2295 | -35.1382 | -30.115
20.2356 | -35.1387 | -30.1219
20.2417 | -35.1392 | -30.1285
20.2478 | -35.1396 | -30.1352
20.2539 | -35.1401 | -30.1422
20.26 | -35.1405 | -30.1488
20.2661 | -35.141 | -30.1558
20.2723 | -35.1415 | -30.1622
20.2784 | -35.142 | -30.1688
20.2845 | -35.1425 | -30.1753
20.2906 | -35.143 | -30.182
20.2967 | -35.1435 | -30.1886
20.3028 | -35.144 | -30.1949
20.3089 | -35.1444 | -30.2018
20.315 | -35.1449 | -30.2084
20.3211 | -35.1454 | -30.215
20.3272 | -35.1459 | -30.2218
20.3333 | -35.1464 | -30.2283
20.3394 | -35.1469 | -30.2348
20.3455 | -35.1474 | -30.2414
20.3516 | -35.1479 | -30.2479
20.3577 | -35.1483 | -30.2543
20.3638 | -35.1488 | -30.261
20.3699 | -35.1493 | -30.2672
20.376 | -35.1498 | -30.2741
20.3821 | -35.1503 | -30.2807
20.3882 | -35.1508 | -30.2876
20.3943 | -35.1513 | -30.2944
20.4004 | -35.1518 | -30.3008
20.4065 | -35.1523 | -30.3074
20.4126 | -35.1527 | -30.3141
20.4188 | -35.1533 | -30.3211
20.4249 | -35.1538 | -30.3277
20.431 | -35.1542 | -30.3347
20.4371 | -35.1547 | -30.3407
20.4432 | -35.1551 | -30.3474
20.4493 | -35.1556 | -30.354
20.4554 | -35.1561 | -30.3605
20.4615 | -35.1566 | -30.3673
20.4676 | -35.157 | -30.3739
20.4737 | -35.1576 | -30.38
20.4798 | -35.1581 | -30.3866
20.4859 | -35.1586 | -30.3931
20.492 | -35.1591 | -30.3999
20.4981 | -35.1596 | -30.4066
20.5042 | -35.1601 | -30.413
20.5103 | -35.1605 | -30.4191
20.5164 | -35.161 | -30.4254
20.5225 | -35.1615 | -30.4322
20.5286 | -35.1619 | -30.439
20.5347 | -35.1624 | -30.4455
20.5408 | -35.1629 | -30.4521
20.5469 | -35.1634 | -30.4584
20.553 | -35.1639 | -30.4651
20.5592 | -35.1644 | -30.4716
20.5653 | -35.1649 | -30.4783
20.5714 | -35.1654 | -30.4845
20.5775 | -35.1658 | -30.4912
20.5836 | -35.1662 | -30.4981
20.5897 | -35.1668 | -30.5045
20.5958 | -35.1673 | -30.5109
20.6019 | -35.1677 | -30.5171
20.608 | -35.1682 | -30.5238
20.6141 | -35.1687 | -30.5304
20.6202 | -35.1692 | -30.5371
20.6263 | -35.1696 | -30.5436
20.6324 | -35.1701 | -30.5504
20.6385 | -35.1705 | -30.557
20.6446 | -35.171 | -30.5636
20.6507 | -35.1715 | -30.5697
20.6568 | -35.172 | -30.5759
20.6629 | -35.1725 | -30.5823
20.669 | -35.173 | -30.5889
20.6751 | -35.1734 | -30.5949
20.6812 | -35.1739 | -30.6011
20.6873 | -35.1744 | -30.6079
20.6934 | -35.1749 | -30.6143
20.6995 | -35.1753 | -30.621
20.7057 | -35.1758 | -30.6275
20.7118 | -35.1763 | -30.6339
20.7179 | -35.1768 | -30.6402
20.724 | -35.1773 | -30.6467
20.7301 | -35.1777 | -30.6535
20.7362 | -35.1782 | -30.6601
20.7423 | -35.1787 | -30.6663
20.7484 | -35.1792 | -30.673
20.7545 | -35.1797 | -30.6793
20.7606 | -35.1802 | -30.6858
20.7667 | -35.1807 | -30.6919
20.7728 | -35.1813 | -30.6985
20.7789 | -35.1818 | -30.7053
20.785 | -35.1822 | -30.7122
20.7911 | -35.1827 | -30.7184
20.7972 | -35.1832 | -30.7249
20.8033 | -35.1836 | -30.7315
20.8094 | -35.1841 | -30.7382
20.8155 | -35.1846 | -30.7449
20.8216 | -35.1851 | -30.7516
20.8277 | -35.1856 | -30.7576
20.8338 | -35.186 | -30.7642
20.8399 | -35.1865 | -30.7706
20.8461 | -35.187 | -30.7768
20.8522 | -35.1875 | -30.7833
20.8583 | -35.188 | -30.7895
20.8644 | -35.1885 | -30.7961
20.8705 | -35.189 | -30.803
20.8766 | -35.1895 | -30.8097
20.8827 | -35.19 | -30.8161
20.8888 | -35.1905 | -30.8227
20.8949 | -35.191 | -30.8286
20.901 | -35.1914 | -30.8353
20.9071 | -35.1919 | -30.8418
20.9132 | -35.1924 | -30.8482
20.9193 | -35.1928 | -30.8544
20.9254 | -35.1933 | -30.8608
20.9315 | -35.1938 | -30.8669
20.9376 | -35.1943 | -30.8735
20.9437 | -35.1948 | -30.8799
20.9498 | -35.1953 | -30.886
20.9559 | -35.1958 | -30.8922
20.962 | -35.1963 | -30.899
20.9681 | -35.1967 | -30.9058
20.9742 | -35.1972 | -30.9123
20.9803 | -35.1977 | -30.9187
20.9864 | -35.1982 | -30.9251
20.9926 | -35.1987 | -30.9317
20.9987 | -35.1991 | -30.9379
21.0048 | -35.1996 | -30.9446
21.0109 | -35.2 | -30.951
21.017 | -35.2005 | -30.957
21.0231 | -35.2011 | -30.9634
21.0292 | -35.2015 | -30.9699
21.0353 | -35.202 | -30.976
21.0414 | -35.2025 | -30.9826
21.0475 | -35.203 | -30.9891
21.0536 | -35.2035 | -30.9951
21.0597 | -35.204 | -31.0012
21.0658 | -35.2044 | -31.0073
21.0719 | -35.2049 | -31.0136
21.078 | -35.2053 | -31.0203
21.0841 | -35.2058 | -31.0263
21.0902 | -35.2063 | -31.0331
21.0963 | -35.2068 | -31.0392
21.1024 | -35.2072 | -31.0453
21.1085 | -35.2077 | -31.0512
21.1146 | -35.2081 | -31.058
21.1207 | -35.2086 | -31.0642
21.1268 | -35.2091 | -31.0707
21.133 | -35.2096 | -31.0771
21.1391 | -35.21 | -31.0831
21.1452 | -35.2105 | -31.0896
21.1513 | -35.211 | -31.0959
21.1574 | -35.2114 | -31.1025
21.1635 | -35.2119 | -31.1086
21.1696 | -35.2124 | -31.1151
21.1757 | -35.2128 | -31.1212
21.1818 | -35.2133 | -31.1278
21.1879 | -35.2138 | -31.134
21.194 | -35.2143 | -31.1404
21.2001 | -35.2148 | -31.1467
21.2062 | -35.2152 | -31.1528
21.2123 | -35.2158 | -31.1591
21.2184 | -35.2162 | -31.1654
21.2245 | -35.2167 | -31.1716
21.2306 | -35.2172 | -31.1779
21.2367 | -35.2176 | -31.1842
21.2428 | -35.2181 | -31.1906
21.2489 | -35.2185 | -31.1968
21.255 | -35.219 | -31.2032
21.2611 | -35.2195 | -31.2094
21.2672 | -35.22 | -31.2157
21.2733 | -35.2204 | -31.2221
21.2795 | -35.2209 | -31.2281
21.2856 | -35.2213 | -31.2347
21.2917 | -35.2218 | -31.2411
21.2978 | -35.2223 | -31.2474
21.3039 | -35.2227 | -31.2535
21.31 | -35.2232 | -31.2597
21.3161 | -35.2237 | -31.266
21.3222 | -35.2242 | -31.2721
21.3283 | -35.2247 | -31.2786
21.3344 | -35.2251 | -31.2849
21.3405 | -35.2255 | -31.2916
21.3466 | -35.226 | -31.2981
21.3527 | -35.2265 | -31.304
21.3588 | -35.227 | -31.3101
21.3649 | -35.2274 | -31.3163
21.371 | -35.2279 | -31.3223
21.3771 | -35.2283 | -31.3285
21.3832 | -35.2288 | -31.3345
21.3893 | -35.2293 | -31.3404
21.3954 | -35.2297 | -31.3466
21.4015 | -35.2302 | -31.3528
21.4076 | -35.2306 | -31.3589
21.4137 | -35.2311 | -31.3652
21.4199 | -35.2315 | -31.3717
21.426 | -35.232 | -31.3778
21.4321 | -35.2325 | -31.3843
21.4382 | -35.233 | -31.3906
21.4443 | -35.2335 | -31.3968
21.4504 | -35.2339 | -31.4032
21.4565 | -35.2344 | -31.4094
21.4626 | -35.2349 | -31.4158
21.4687 | -35.2353 | -31.4223
21.4748 | -35.2358 | -31.4285
21.4809 | -35.2362 | -31.4346
21.487 | -35.2367 | -31.4407
21.4931 | -35.2371 | -31.4466
21.4992 | -35.2376 | -31.4528
21.5053 | -35.2381 | -31.4593
21.5114 | -35.2385 | -31.4653
21.5175 | -35.2389 | -31.4718
21.5236 | -35.2394 | -31.4781
21.5297 | -35.2398 | -31.484
21.5358 | -35.2403 | -31.4901
21.5419 | -35.2407 | -31.4971
21.548 | -35.2411 | -31.5034
21.5541 | -35.2416 | -31.5096
21.5602 | -35.2421 | -31.5161
21.5664 | -35.2425 | -31.5222
21.5725 | -35.243 | -31.5284
21.5786 | -35.2434 | -31.5348
21.5847 | -35.2438 | -31.5412
21.5908 | -35.2443 | -31.5474
21.5969 | -35.2447 | -31.5538
21.603 | -35.2451 | -31.5602
21.6091 | -35.2456 | -31.5668
21.6152 | -35.2461 | -31.5729
21.6213 | -35.2466 | -31.5793
21.6274 | -35.2471 | -31.5855
21.6335 | -35.2475 | -31.5912
21.6396 | -35.2479 | -31.5978
21.6457 | -35.2484 | -31.6043
21.6518 | -35.2488 | -31.6101
21.6579 | -35.2493 | -31.6159
21.664 | -35.2498 | -31.6221
21.6701 | -35.2503 | -31.6276
21.6762 | -35.2507 | -31.6338
21.6823 | -35.2512 | -31.6402
21.6884 | -35.2516 | -31.6467
21.6945 | -35.2521 | -31.6528
21.7006 | -35.2525 | -31.6591
21.7068 | -35.2529 | -31.6651
21.7129 | -35.2534 | -31.6717
21.719 | -35.2539 | -31.6783
21.7251 | -35.2543 | -31.6845
21.7312 | -35.2547 | -31.6906
21.7373 | -35.2552 | -31.6966
21.7434 | -35.2557 | -31.7029
21.7495 | -35.2561 | -31.7092
21.7556 | -35.2565 | -31.7157
21.7617 | -35.257 | -31.7217
21.7678 | -35.2575 | -31.7281
21.7739 | -35.258 | -31.734
21.78 | -35.2584 | -31.7399
21.7861 | -35.2589 | -31.7461
21.7922 | -35.2593 | -31.7524
21.7983 | -35.2598 | -31.7586
21.8044 | -35.2603 | -31.7651
21.8105 | -35.2607 | -31.7712
21.8166 | -35.2612 | -31.7775
21.8227 | -35.2617 | -31.7836
21.8288 | -35.2621 | -31.7898
21.8349 | -35.2625 | -31.7959
21.841 | -35.263 | -31.8022
21.8471 | -35.2634 | -31.8082
21.8533 | -35.2638 | -31.8143
21.8594 | -35.2643 | -31.8204
21.8655 | -35.2648 | -31.8268
21.8716 | -35.2653 | -31.8332
21.8777 | -35.2657 | -31.8389
21.8838 | -35.2662 | -31.845
21.8899 | -35.2666 | -31.8511
21.896 | -35.2671 | -31.8575
21.9021 | -35.2675 | -31.8639
21.9082 | -35.268 | -31.8701
21.9143 | -35.2684 | -31.8764
21.9204 | -35.2688 | -31.8825
21.9265 | -35.2693 | -31.8885
21.9326 | -35.2697 | -31.8948
21.9387 | -35.2702 | -31.9009
21.9448 | -35.2706 | -31.907
21.9509 | -35.2712 | -31.9132
21.957 | -35.2716 | -31.9196
21.9631 | -35.2721 | -31.9258
21.9692 | -35.2725 | -31.9322
21.9753 | -35.273 | -31.9382
21.9814 | -35.2734 | -31.9445
21.9875 | -35.2738 | -31.9505
21.9937 | -35.2742 | -31.9568
21.9998 | -35.2747 | -31.9631
22.0059 | -35.2752 | -31.9694
22.012 | -35.2756 | -31.9753
22.0181 | -35.2761 | -31.9813
22.0242 | -35.2765 | -31.9875
22.0303 | -35.277 | -31.9941
22.0364 | -35.2774 | -32.0002
22.0425 | -35.2779 | -32.0064
22.0486 | -35.2784 | -32.0124
22.0547 | -35.2788 | -32.018
22.0608 | -35.2792 | -32.0239
22.0669 | -35.2797 | -32.0298
22.073 | -35.2801 | -32.0357
22.0791 | -35.2806 | -32.0422
22.0852 | -35.281 | -32.0482
22.0913 | -35.2814 | -32.0541
22.0974 | -35.2818 | -32.0602
22.1035 | -35.2823 | -32.066
22.1096 | -35.2827 | -32.0722
22.1157 | -35.2831 | -32.0783
22.1218 | -35.2836 | -32.0846
22.1279 | -35.2841 | -32.0904
22.134 | -35.2845 | -32.0965
22.1402 | -35.285 | -32.1025
22.1463 | -35.2854 | -32.1085
22.1524 | -35.2859 | -32.1149
22.1585 | -35.2863 | -32.1212
22.1646 | -35.2868 | -32.1269
22.1707 | -35.2873 | -32.1333
22.1768 | -35.2878 | -32.1393
22.1829 | -35.2882 | -32.1456
22.189 | -35.2886 | -32.1517
22.1951 | -35.2891 | -32.1578
22.2012 | -35.2895 | -32.1642
22.2073 | -35.29 | -32.1705
22.2134 | -35.2904 | -32.1764
22.2195 | -35.2909 | -32.1826
22.2256 | -35.2913 | -32.1889
22.2317 | -35.2918 | -32.1948
22.2378 | -35.2922 | -32.201
22.2439 | -35.2927 | -32.2073
22.25 | -35.2931 | -32.2135
22.2561 | -35.2936 | -32.2196
22.2622 | -35.294 | -32.226
22.2683 | -35.2944 | -32.2317
22.2744 | -35.2948 | -32.2378
22.2806 | -35.2953 | -32.2439
22.2867 | -35.2957 | -32.2499
22.2928 | -35.2961 | -32.256
22.2989 | -35.2966 | -32.262
22.305 | -35.297 | -32.2677
22.3111 | -35.2974 | -32.2739
22.3172 | -35.2978 | -32.28
22.3233 | -35.2983 | -32.2862
22.3294 | -35.2987 | -32.2922
22.3355 | -35.2992 | -32.2982
22.3416 | -35.2997 | -32.3045
22.3477 | -35.3002 | -32.3107
22.3538 | -35.3006 | -32.3168
22.3599 | -35.301 | -32.3227
22.366 | -35.3015 | -32.3288
22.3721 | -35.302 | -32.3347
22.3782 | -35.3024 | -32.3409
22.3843 | -35.3028 | -32.347
22.3904 | -35.3033 | -32.3534
22.3965 | -35.3038 | -32.3594
22.4026 | -35.3042 | -32.3652
22.4087 | -35.3046 | -32.3712
22.4148 | -35.305 | -32.3772
22.4209 | -35.3055 | -32.3833
22.4271 | -35.3059 | -32.3892
22.4332 | -35.3063 | -32.3951
22.4393 | -35.3068 | -32.401
22.4454 | -35.3072 | -32.4066
22.4515 | -35.3076 | -32.4127
22.4576 | -35.3081 | -32.4188
22.4637 | -35.3085 | -32.425
22.4698 | -35.3089 | -32.431
22.4759 | -35.3094 | -32.4368
22.482 | -35.3098 | -32.4428
22.4881 | -35.3102 | -32.4489
22.4942 | -35.3107 | -32.455
22.5003 | -35.3111 | -32.461
22.5064 | -35.3115 | -32.4668
22.5125 | -35.3119 | -32.4726
22.5186 | -35.3123 | -32.4786
22.5247 | -35.3128 | -32.4847
22.5308 | -35.3132 | -32.4908
22.5369 | -35.3137 | -32.4969
22.543 | -35.3141 | -32.5031
22.5491 | -35.3145 | -32.5088
22.5552 | -35.3149 | -32.5151
22.5613 | -35.3153 | -32.5213
22.5675 | -35.3158 | -32.5275
22.5736 | -35.3162 | -32.5336
22.5797 | -35.3166 | -32.5396
22.5858 | -35.317 | -32.5455
22.5919 | -35.3175 | -32.5514
22.598 | -35.3179 | -32.5576
22.6041 | -35.3183 | -32.5637
22.6102 | -35.3188 | -32.5696
22.6163 | -35.3192 | -32.5754
22.6224 | -35.3196 | -32.5815
22.6285 | -35.3201 | -32.5876
22.6346 | -35.3205 | -32.5934
22.6407 | -35.3209 | -32.5996
22.6468 | -35.3214 | -32.6057
22.6529 | -35.3218 | -32.6117
22.659 | -35.3222 | -32.6177
22.6651 | -35.3226 | -32.6237
22.6712 | -35.323 | -32.6296
22.6773 | -35.3235 | -32.6357
22.6834 | -35.3239 | -32.642
22.6895 | -35.3244 | -32.6478
22.6956 | -35.3248 | -32.6538
22.7017 | -35.3253 | -32.6599
22.7079 | -35.3257 | -32.6661
22.714 | -35.3262 | -32.6721
22.7201 | -35.3266 | -32.6782
22.7262 | -35.3271 | -32.6841
22.7323 | -35.3275 | -32.6902
22.7384 | -35.3279 | -32.6962
22.7445 | -35.3283 | -32.7024
22.7506 | -35.3288 | -32.7084
22.7567 | -35.3292 | -32.7145
22.7628 | -35.3296 | -32.7205
22.7689 | -35.3301 | -32.7265
22.775 | -35.3305 | -32.7329
22.7811 | -35.3309 | -32.739
22.7872 | -35.3313 | -32.745
22.7933 | -35.3318 | -32.7507
22.7994 | -35.3323 | -32.7568
22.8055 | -35.3326 | -32.7629
22.8116 | -35.3331 | -32.7689
22.8177 | -35.3335 | -32.7751
22.8238 | -35.334 | -32.7813
22.8299 | -35.3344 | -32.787
22.836 | -35.3348 | -32.7933
22.8421 | -35.3352 | -32.7992
22.8482 | -35.3356 | -32.8052
22.8544 | -35.336 | -32.8114
22.8605 | -35.3365 | -32.8172
22.8666 | -35.3369 | -32.8234
22.8727 | -35.3373 | -32.8293
22.8788 | -35.3378 | -32.8353
22.8849 | -35.3382 | -32.8413
22.891 | -35.3387 | -32.8473
22.8971 | -35.3391 | -32.8534
22.9032 | -35.3395 | -32.8593
22.9093 | -35.3399 | -32.8653
22.9154 | -35.3404 | -32.8713
22.9215 | -35.3408 | -32.8772
22.9276 | -35.3412 | -32.8834
22.9337 | -35.3416 | -32.8891
22.9398 | -35.3421 | -32.8951
22.9459 | -35.3425 | -32.9012
22.952 | -35.3429 | -32.9073
22.9581 | -35.3434 | -32.9135
22.9642 | -35.3438 | -32.9196
22.9703 | -35.3442 | -32.9254
22.9764 | -35.3447 | -32.9313
22.9825 | -35.3451 | -32.9373
22.9886 | -35.3455 | -32.9433
22.9948 | -35.3459 | -32.9495
23.0009 | -35.3463 | -32.9556
23.007 | -35.3468 | -32.9617
23.0131 | -35.3472 | -32.9676
23.0192 | -35.3477 | -32.9734
23.0253 | -35.3481 | -32.9796
23.0314 | -35.3486 | -32.9855
23.0375 | -35.349 | -32.9916
23.0436 | -35.3494 | -32.9974
23.0497 | -35.3499 | -33.0035
23.0558 | -35.3503 | -33.0094
23.0619 | -35.3508 | -33.0154
23.068 | -35.3512 | -33.0215
23.0741 | -35.3516 | -33.0268
23.0802 | -35.352 | -33.0328
23.0863 | -35.3524 | -33.0389
23.0924 | -35.3528 | -33.0448
23.0985 | -35.3532 | -33.0509
23.1046 | -35.3537 | -33.0568
23.1107 | -35.3541 | -33.0626
23.1168 | -35.3545 | -33.0687
23.1229 | -35.3549 | -33.0743
23.129 | -35.3553 | -33.0805
23.1351 | -35.3557 | -33.0866
23.1413 | -35.3561 | -33.0923
23.1474 | -35.3566 | -33.0982
23.1535 | -35.357 | -33.104
23.1596 | -35.3574 | -33.1097
23.1657 | -35.3578 | -33.1153
23.1718 | -35.3583 | -33.1214
23.1779 | -35.3587 | -33.1274
23.184 | -35.3592 | -33.133
23.1901 | -35.3596 | -33.139
23.1962 | -35.36 | -33.1448
23.2023 | -35.3604 | -33.1507
23.2084 | -35.3609 | -33.1566
23.2145 | -35.3613 | -33.1625
23.2206 | -35.3618 | -33.1685
23.2267 | -35.3622 | -33.1745
23.2328 | -35.3626 | -33.1803
23.2389 | -35.363 | -33.1863
23.245 | -35.3635 | -33.1926
23.2511 | -35.3639 | -33.198
23.2572 | -35.3643 | -33.204
23.2633 | -35.3648 | -33.21
23.2694 | -35.3652 | -33.2159
23.2755 | -35.3656 | -33.2219
23.2817 | -35.366 | -33.2279
23.2878 | -35.3664 | -33.2341
23.2939 | -35.3668 | -33.2399
23.3 | -35.3673 | -33.2458
23.3061 | -35.3677 | -33.2518
23.3122 | -35.3681 | -33.2578
23.3183 | -35.3685 | -33.2635
23.3244 | -35.3689 | -33.2696
23.3305 | -35.3693 | -33.2758
23.3366 | -35.3697 | -33.2819
23.3427 | -35.3701 | -33.2876
23.3488 | -35.3706 | -33.2936
23.3549 | -35.371 | -33.2996
23.361 | -35.3714 | -33.3052
23.3671 | -35.3719 | -33.3113
23.3732 | -35.3723 | -33.3172
23.3793 | -35.3727 | -33.3229
23.3854 | -35.3732 | -33.3292
23.3915 | -35.3735 | -33.3351
23.3976 | -35.374 | -33.341
23.4037 | -35.3744 | -33.3467
23.4098 | -35.3749 | -33.3527
23.4159 | -35.3753 | -33.3588
23.422 | -35.3757 | -33.3647
23.4282 | -35.3761 | -33.3705
23.4343 | -35.3765 | -33.3763
23.4404 | -35.377 | -33.3827
23.4465 | -35.3773 | -33.3886
23.4526 | -35.3778 | -33.3949
23.4587 | -35.3781 | -33.4008
23.4648 | -35.3785 | -33.4065
23.4709 | -35.3789 | -33.4125
23.477 | -35.3793 | -33.4184
23.4831 | -35.3797 | -33.4242
23.4892 | -35.3802 | -33.4302
23.4953 | -35.3806 | -33.436
23.5014 | -35.381 | -33.4418
23.5075 | -35.3813 | -33.4475
23.5136 | -35.3818 | -33.4531
23.5197 | -35.3822 | -33.4591
23.5258 | -35.3826 | -33.4649
23.5319 | -35.383 | -33.4706
23.538 | -35.3834 | -33.4767
23.5441 | -35.3838 | -33.4823
23.5502 | -35.3843 | -33.4886
23.5563 | -35.3847 | -33.4944
23.5624 | -35.3851 | -33.5005
23.5686 | -35.3855 | -33.5064
23.5747 | -35.3859 | -33.5119
23.5808 | -35.3863 | -33.5179
23.5869 | -35.3868 | -33.5236
23.593 | -35.3872 | -33.5296
23.5991 | -35.3876 | -33.5356
23.6052 | -35.388 | -33.5412
23.6113 | -35.3885 | -33.5472
23.6174 | -35.3889 | -33.5529
23.6235 | -35.3893 | -33.5591
23.6296 | -35.3897 | -33.5651
23.6357 | -35.3901 | -33.5714
23.6418 | -35.3906 | -33.5769
23.6479 | -35.391 | -33.5827
23.654 | -35.3914 | -33.5888
23.6601 | -35.3918 | -33.5946
23.6662 | -35.3922 | -33.6006
23.6723 | -35.3927 | -33.6064
23.6784 | -35.3931 | -33.6122
23.6845 | -35.3935 | -33.6178
23.6906 | -35.3939 | -33.6238
23.6967 | -35.3943 | -33.6299
23.7028 | -35.3947 | -33.636
23.7089 | -35.3951 | -33.6422
23.7151 | -35.3955 | -33.6481
23.7212 | -35.3959 | -33.654
23.7273 | -35.3963 | -33.6597
23.7334 | -35.3967 | -33.6657
23.7395 | -35.3972 | -33.6716
23.7456 | -35.3975 | -33.6775
23.7517 | -35.398 | -33.6835
23.7578 | -35.3984 | -33.6894
23.7639 | -35.3987 | -33.6951
23.77 | -35.3992 | -33.7009
23.7761 | -35.3996 | -33.7066
23.7822 | -35.3999 | -33.7126
23.7883 | -35.4003 | -33.7185
23.7944 | -35.4007 | -33.7245
23.8005 | -35.4012 | -33.7306
23.8066 | -35.4016 | -33.7365
23.8127 | -35.4021 | -33.7427
23.8188 | -35.4025 | -33.7488
23.8249 | -35.4029 | -33.7549
23.831 | -35.4033 | -33.761
23.8371 | -35.4037 | -33.7667
23.8432 | -35.4041 | -33.7725
23.8493 | -35.4046 | -33.7781
23.8555 | -35.405 | -33.7842
23.8616 | -35.4054 | -33.7904
23.8677 | -35.4058 | -33.7963
23.8738 | -35.4062 | -33.8023
23.8799 | -35.4067 | -33.8084
23.886 | -35.4071 | -33.814
23.8921 | -35.4075 | -33.8199
23.8982 | -35.4079 | -33.8257
23.9043 | -35.4083 | -33.8319
23.9104 | -35.4086 | -33.8378
23.9165 | -35.409 | -33.8435
23.9226 | -35.4094 | -33.8495
23.9287 | -35.4098 | -33.8556
23.9348 | -35.4102 | -33.8613
23.9409 | -35.4106 | -33.8672
23.947 | -35.4111 | -33.8729
23.9531 | -35.4115 | -33.8788
23.9592 | -35.4119 | -33.8846
23.9653 | -35.4124 | -33.8905
23.9714 | -35.4128 | -33.8964
23.9775 | -35.4131 | -33.9021
23.9836 | -35.4136 | -33.9081
23.9897 | -35.414 | -33.9137
23.9958 | -35.4145 | -33.9197
24.002 | -35.4149 | -33.9256
24.0081 | -35.4153 | -33.9313
24.0142 | -35.4157 | -33.9372
24.0203 | -35.4161 | -33.9427
24.0264 | -35.4165 | -33.9485
24.0325 | -35.4169 | -33.9545
24.0386 | -35.4173 | -33.9604
24.0447 | -35.4176 | -33.9658
24.0508 | -35.4181 | -33.9714
24.0569 | -35.4185 | -33.9774
24.063 | -35.4189 | -33.9832
24.0691 | -35.4193 | -33.9889
24.0752 | -35.4197 | -33.995
24.0813 | -35.4201 | -34.0008
24.0874 | -35.4206 | -34.0063
24.0935 | -35.421 | -34.0122
24.0996 | -35.4213 | -34.0178
24.1057 | -35.4217 | -34.0234
24.1118 | -35.4221 | -34.0295
24.1179 | -35.4224 | -34.0354
24.124 | -35.4229 | -34.0412
24.1301 | -35.4233 | -34.0469
24.1362 | -35.4237 | -34.0526
24.1424 | -35.4241 | -34.0585
24.1485 | -35.4245 | -34.0642
24.1546 | -35.4249 | -34.0701
24.1607 | -35.4253 | -34.0758
24.1668 | -35.4258 | -34.0815
24.1729 | -35.4262 | -34.0874
24.179 | -35.4265 | -34.0933
24.1851 | -35.427 | -34.0992
24.1912 | -35.4274 | -34.1052
24.1973 | -35.4277 | -34.1112
24.2034 | -35.4282 | -34.1168
24.2095 | -35.4286 | -34.1228
24.2156 | -35.429 | -34.1288
24.2217 | -35.4294 | -34.1347
24.2278 | -35.4299 | -34.1405
24.2339 | -35.4303 | -34.1461
24.24 | -35.4306 | -34.1522
24.2461 | -35.4311 | -34.1579
24.2522 | -35.4315 | -34.1637
24.2583 | -35.4319 | -34.1697
24.2644 | -35.4323 | -34.1756
24.2705 | -35.4327 | -34.1815
24.2766 | -35.4332 | -34.1872
24.2827 | -35.4336 | -34.1928
24.2889 | -35.4339 | -34.1987
24.295 | -35.4343 | -34.2046
24.3011 | -35.4347 | -34.2106
24.3072 | -35.4352 | -34.2161
24.3133 | -35.4356 | -34.2217
24.3194 | -35.436 | -34.2277
24.3255 | -35.4364 | -34.2337
24.3316 | -35.4368 | -34.2397
24.3377 | -35.4372 | -34.2456
24.3438 | -35.4375 | -34.2519
24.3499 | -35.4379 | -34.2578
24.356 | -35.4383 | -34.2635
24.3621 | -35.4386 | -34.2692
24.3682 | -35.439 | -34.2751
24.3743 | -35.4395 | -34.2809
24.3804 | -35.4398 | -34.2864
24.3865 | -35.4403 | -34.2922
24.3926 | -35.4407 | -34.2981
24.3987 | -35.4411 | -34.3038
24.4048 | -35.4415 | -34.3095
24.4109 | -35.4419 | -34.315
24.417 | -35.4423 | -34.3208
24.4231 | -35.4427 | -34.3269
24.4293 | -35.4431 | -34.3327
24.4354 | -35.4435 | -34.3382
24.4415 | -35.4439 | -34.344
24.4476 | -35.4443 | -34.3493
24.4537 | -35.4447 | -34.3551
24.4598 | -35.4452 | -34.3609
24.4659 | -35.4456 | -34.3667
24.472 | -35.4459 | -34.3726
24.4781 | -35.4463 | -34.3784
24.4842 | -35.4467 | -34.384
24.4903 | -35.447 | -34.3896
24.4964 | -35.4474 | -34.3951
24.5025 | -35.4479 | -34.401
24.5086 | -35.4483 | -34.4068
24.5147 | -35.4487 | -34.4125
24.5208 | -35.4491 | -34.4177
24.5269 | -35.4495 | -34.4234
24.533 | -35.4499 | -34.429
24.5391 | -35.4503 | -34.4348
24.5452 | -35.4507 | -34.4404
24.5513 | -35.4511 | -34.4463
24.5574 | -35.4516 | -34.452
24.5635 | -35.452 | -34.4577
24.5696 | -35.4524 | -34.4633
24.5758 | -35.4528 | -34.469
24.5819 | -35.4532 | -34.4748
24.588 | -35.4536 | -34.4808
24.5941 | -35.454 | -34.4864
24.6002 | -35.4544 | -34.4923
24.6063 | -35.4548 | -34.4982
24.6124 | -35.4552 | -34.5043
24.6185 | -35.4556 | -34.5101
24.6246 | -35.456 | -34.5159
24.6307 | -35.4564 | -34.5218
24.6368 | -35.4568 | -34.5276
24.6429 | -35.4572 | -34.5331
24.649 | -35.4575 | -34.5386
24.6551 | -35.4579 | -34.5445
24.6612 | -35.4584 | -34.5501
24.6673 | -35.4588 | -34.5558
24.6734 | -35.4592 | -34.5615
24.6795 | -35.4596 | -34.5674
24.6856 | -35.46 | -34.5733
24.6917 | -35.4604 | -34.5791
24.6978 | -35.4608 | -34.5848
24.7039 | -35.4612 | -34.5906
24.71 | -35.4616 | -34.5963
24.7162 | -35.462 | -34.602
24.7223 | -35.4624 | -34.6074
24.7284 | -35.4628 | -34.6128
24.7345 | -35.4632 | -34.6187
24.7406 | -35.4636 | -34.6243
24.7467 | -35.464 | -34.6301
24.7528 | -35.4644 | -34.6359
24.7589 | -35.4647 | -34.6416
24.765 | -35.4651 | -34.6469
24.7711 | -35.4655 | -34.6526
24.7772 | -35.4659 | -34.6583
24.7833 | -35.4663 | -34.6639
24.7894 | -35.4667 | -34.6695
24.7955 | -35.4671 | -34.6753
24.8016 | -35.4675 | -34.6809
24.8077 | -35.4679 | -34.6869
24.8138 | -35.4683 | -34.6927
24.8199 | -35.4687 | -34.6985
24.826 | -35.4691 | -34.7042
24.8321 | -35.4695 | -34.7099
24.8382 | -35.4699 | -34.7154
24.8443 | -35.4703 | -34.721
24.8504 | -35.4707 | -34.7268
24.8565 | -35.4711 | -34.7324
24.8627 | -35.4715 | -34.7381
24.8688 | -35.4719 | -34.7439
24.8749 | -35.4723 | -34.7495
24.881 | -35.4727 | -34.7554
24.8871 | -35.4731 | -34.7609
24.8932 | -35.4736 | -34.7663
24.8993 | -35.474 | -34.7717
24.9054 | -35.4744 | -34.7774
24.9115 | -35.4749 | -34.7832
24.9176 | -35.4753 | -34.7889
24.9237 | -35.4757 | -34.7943
24.9298 | -35.476 | -34.7999
24.9359 | -35.4765 | -34.8058
24.942 | -35.4769 | -34.811
24.9481 | -35.4773 | -34.8167
24.9542 | -35.4777 | -34.8225
24.9603 | -35.4781 | -34.8284
24.9664 | -35.4785 | -34.834
24.9725 | -35.4789 | -34.8398
24.9786 | -35.4792 | -34.8453
24.9847 | -35.4796 | -34.8509
24.9908 | -35.48 | -34.8565
24.9969 | -35.4804 | -34.8624
25.0031 | -35.4808 | -34.8681
25.0092 | -35.4812 | -34.8736
25.0153 | -35.4816 | -34.8791
25.0214 | -35.4819 | -34.8851
25.0275 | -35.4823 | -34.8907
25.0336 | -35.4828 | -34.8963
25.0397 | -35.4831 | -34.9022
25.0458 | -35.4835 | -34.9078
25.0519 | -35.4839 | -34.9137
25.058 | -35.4843 | -34.9194
25.0641 | -35.4847 | -34.9249
25.0702 | -35.4851 | -34.9307
25.0763 | -35.4855 | -34.9363
25.0824 | -35.4859 | -34.9419
25.0885 | -35.4862 | -34.9477
25.0946 | -35.4866 | -34.9533
25.1007 | -35.487 | -34.959
25.1068 | -35.4874 | -34.9644
25.1129 | -35.4878 | -34.9702
25.119 | -35.4882 | -34.9757
25.1251 | -35.4886 | -34.9813
25.1312 | -35.489 | -34.9868
25.1373 | -35.4894 | -34.9926
25.1435 | -35.4898 | -34.9985
25.1496 | -35.4902 | -35.0038
25.1557 | -35.4906 | -35.0095
25.1618 | -35.491 | -35.0154
25.1679 | -35.4914 | -35.021
25.174 | -35.4918 | -35.0267
25.1801 | -35.4922 | -35.0325
25.1862 | -35.4926 | -35.0382
25.1923 | -35.493 | -35.0437
25.1984 | -35.4934 | -35.0493
25.2045 | -35.4938 | -35.0549
25.2106 | -35.4941 | -35.0605
25.2167 | -35.4945 | -35.0662
25.2228 | -35.4949 | -35.0717
25.2289 | -35.4953 | -35.0773
25.235 | -35.4957 | -35.0829
25.2411 | -35.4961 | -35.0885
25.2472 | -35.4965 | -35.0942
25.2533 | -35.4969 | -35.0999
25.2594 | -35.4972 | -35.1053
25.2655 | -35.4976 | -35.1108
25.2716 | -35.498 | -35.1165
25.2777 | -35.4984 | -35.1221
25.2838 | -35.4988 | -35.1278
25.29 | -35.4992 | -35.1334
25.2961 | -35.4995 | -35.139
25.3022 | -35.4999 | -35.1444
25.3083 | -35.5003 | -35.1501
25.3144 | -35.5006 | -35.1557
25.3205 | -35.501 | -35.1613
25.3266 | -35.5014 | -35.1669
25.3327 | -35.5018 | -35.1726
25.3388 | -35.5022 | -35.1783
25.3449 | -35.5026 | -35.184
25.351 | -35.5029 | -35.1896
25.3571 | -35.5033 | -35.1954
25.3632 | -35.5037 | -35.2008
25.3693 | -35.5041 | -35.2065
25.3754 | -35.5045 | -35.2122
25.3815 | -35.5049 | -35.2179
25.3876 | -35.5053 | -35.2236
25.3937 | -35.5056 | -35.2291
25.3998 | -35.5061 | -35.2343
25.4059 | -35.5064 | -35.2401
25.412 | -35.5068 | -35.2456
25.4181 | -35.5072 | -35.2512
25.4242 | -35.5076 | -35.2567
25.4304 | -35.508 | -35.2622
25.4365 | -35.5084 | -35.2676
25.4426 | -35.5088 | -35.2732
25.4487 | -35.5092 | -35.2789
25.4548 | -35.5096 | -35.2846
25.4609 | -35.51 | -35.2903
25.467 | -35.5104 | -35.2954
25.4731 | -35.5109 | -35.301
25.4792 | -35.5113 | -35.3065
25.4853 | -35.5117 | -35.3119
25.4914 | -35.5121 | -35.3175
25.4975 | -35.5125 | -35.3226
25.5036 | -35.5129 | -35.3284
25.5097 | -35.5133 | -35.3339
25.5158 | -35.5137 | -35.3395
25.5219 | -35.5141 | -35.345
25.528 | -35.5144 | -35.3505
25.5341 | -35.5149 | -35.3559
25.5402 | -35.5152 | -35.3616
25.5463 | -35.5156 | -35.3672
25.5524 | -35.516 | -35.3728
25.5585 | -35.5164 | -35.3781
25.5646 | -35.5168 | -35.3838
25.5707 | -35.5171 | -35.3895
25.5769 | -35.5175 | -35.3949
25.583 | -35.5179 | -35.4007
25.5891 | -35.5183 | -35.4062
25.5952 | -35.5187 | -35.4117
25.6013 | -35.5191 | -35.4169
25.6074 | -35.5196 | -35.4224
25.6135 | -35.52 | -35.4279
25.6196 | -35.5204 | -35.4338
25.6257 | -35.5208 | -35.4393
25.6318 | -35.5211 | -35.4448
25.6379 | -35.5215 | -35.4504
25.644 | -35.5219 | -35.4558
25.6501 | -35.5223 | -35.4611
25.6562 | -35.5227 | -35.4665
25.6623 | -35.5231 | -35.4718
25.6684 | -35.5235 | -35.4774
25.6745 | -35.5238 | -35.4828
25.6806 | -35.5242 | -35.4881
25.6867 | -35.5246 | -35.4937
25.6928 | -35.525 | -35.4993
25.6989 | -35.5253 | -35.5049
25.705 | -35.5257 | -35.5107
25.7111 | -35.5261 | -35.5163
25.7173 | -35.5265 | -35.5214
25.7234 | -35.5268 | -35.5269
25.7295 | -35.5272 | -35.5325
25.7356 | -35.5276 | -35.538
25.7417 | -35.528 | -35.5434
25.7478 | -35.5283 | -35.5487
25.7539 | -35.5287 | -35.554
25.76 | -35.5291 | -35.5595
25.7661 | -35.5295 | -35.5648
25.7722 | -35.5298 | -35.5704
25.7783 | -35.5302 | -35.576
25.7844 | -35.5306 | -35.5812
25.7905 | -35.531 | -35.587
25.7966 | -35.5313 | -35.5924
25.8027 | -35.5317 | -35.5978
25.8088 | -35.532 | -35.6034
25.8149 | -35.5324 | -35.609
25.821 | -35.5327 | -35.6142
25.8271 | -35.5332 | -35.6199
25.8332 | -35.5336 | -35.6256
25.8393 | -35.5339 | -35.631
25.8454 | -35.5343 | -35.6362
25.8515 | -35.5347 | -35.6417
25.8576 | -35.5351 | -35.6471
25.8638 | -35.5355 | -35.6527
25.8699 | -35.5359 | -35.6585
25.876 | -35.5363 | -35.6641
25.8821 | -35.5366 | -35.6699
25.8882 | -35.537 | -35.6752
25.8943 | -35.5375 | -35.6809
25.9004 | -35.5378 | -35.6863
25.9065 | -35.5382 | -35.692
25.9126 | -35.5386 | -35.6975
25.9187 | -35.539 | -35.703
25.9248 | -35.5394 | -35.7086
25.9309 | -35.5398 | -35.714
25.937 | -35.5402 | -35.7197
25.9431 | -35.5405 | -35.7252
25.9492 | -35.5409 | -35.7305
25.9553 | -35.5413 | -35.7359
25.9614 | -35.5417 | -35.741
25.9675 | -35.5421 | -35.7464
25.9736 | -35.5425 | -35.7517
25.9797 | -35.5429 | -35.7576
25.9858 | -35.5433 | -35.7628
25.9919 | -35.5436 | -35.7684
25.998 | -35.544 | -35.7737
26.0042 | -35.5443 | -35.7793
26.0103 | -35.5447 | -35.7846
26.0164 | -35.5451 | -35.79
26.0225 | -35.5455 | -35.7955
26.0286 | -35.5459 | -35.8009
26.0347 | -35.5463 | -35.8064
26.0408 | -35.5466 | -35.8118
26.0469 | -35.547 | -35.8172
26.053 | -35.5474 | -35.8227
26.0591 | -35.5478 | -35.8282
26.0652 | -35.5482 | -35.8334
26.0713 | -35.5486 | -35.839
26.0774 | -35.549 | -35.8444
26.0835 | -35.5493 | -35.8499
26.0896 | -35.5497 | -35.8553
26.0957 | -35.5501 | -35.8606
26.1018 | -35.5505 | -35.866
26.1079 | -35.5508 | -35.8713
26.114 | -35.5512 | -35.8767
26.1201 | -35.5515 | -35.882
26.1262 | -35.5518 | -35.8877
26.1323 | -35.5522 | -35.8928
26.1384 | -35.5526 | -35.8985
26.1445 | -35.553 | -35.9039
26.1507 | -35.5533 | -35.9092
26.1568 | -35.5537 | -35.9147
26.1629 | -35.554 | -35.9203
26.169 | -35.5544 | -35.9257
26.1751 | -35.5548 | -35.9313
26.1812 | -35.5552 | -35.9367
26.1873 | -35.5556 | -35.9421
26.1934 | -35.556 | -35.9475
26.1995 | -35.5564 | -35.9529
26.2056 | -35.5566 | -35.958
26.2117 | -35.557 | -35.9635
26.2178 | -35.5574 | -35.969
26.2239 | -35.5578 | -35.9746
26.23 | -35.5581 | -35.9802
26.2361 | -35.5585 | -35.9857
26.2422 | -35.5588 | -35.9911
26.2483 | -35.5592 | -35.9963
26.2544 | -35.5596 | -36.0015
26.2605 | -35.56 | -36.007
26.2666 | -35.5603 | -36.0122
26.2727 | -35.5607 | -36.0177
26.2788 | -35.561 | -36.0229
26.2849 | -35.5614 | -36.0282
26.2911 | -35.5617 | -36.0335
26.2972 | -35.5621 | -36.0389
26.3033 | -35.5625 | -36.0443
26.3094 | -35.5628 | -36.0501
26.3155 | -35.5632 | -36.0552
26.3216 | -35.5635 | -36.0607
26.3277 | -35.5638 | -36.066
26.3338 | -35.5642 | -36.0713
26.3399 | -35.5645 | -36.0767
26.346 | -35.565 | -36.0819
26.3521 | -35.5653 | -36.0877
26.3582 | -35.5657 | -36.0931
26.3643 | -35.566 | -36.0986
26.3704 | -35.5664 | -36.1038
26.3765 | -35.5668 | -36.1091
26.3826 | -35.5672 | -36.1144
26.3887 | -35.5676 | -36.1197
26.3948 | -35.5679 | -36.1249
26.4009 | -35.5682 | -36.1304
26.407 | -35.5686 | -36.136
26.4131 | -35.5689 | -36.1412
26.4192 | -35.5693 | -36.1465
26.4253 | -35.5697 | -36.1519
26.4314 | -35.57 | -36.1572
26.4376 | -35.5704 | -36.1625
26.4437 | -35.5707 | -36.1678
26.4498 | -35.5711 | -36.1733
26.4559 | -35.5715 | -36.1781
26.462 | -35.5718 | -36.1837
26.4681 | -35.5722 | -36.1891
26.4742 | -35.5726 | -36.1948
26.4803 | -35.573 | -36.2001
26.4864 | -35.5733 | -36.2055
26.4925 | -35.5737 | -36.2107
26.4986 | -35.574 | -36.2162
26.5047 | -35.5744 | -36.2219
26.5108 | -35.5747 | -36.2272
26.5169 | -35.5751 | -36.2326
26.523 | -35.5754 | -36.2383
26.5291 | -35.5758 | -36.2435
26.5352 | -35.5761 | -36.2489
26.5413 | -35.5765 | -36.2543
26.5474 | -35.5769 | -36.2595
26.5535 | -35.5772 | -36.265
26.5596 | -35.5775 | -36.2701
26.5657 | -35.5779 | -36.2754
26.5718 | -35.5783 | -36.281
26.578 | -35.5786 | -36.2864
26.5841 | -35.579 | -36.2919
26.5902 | -35.5794 | -36.2969
26.5963 | -35.5798 | -36.3025
26.6024 | -35.5801 | -36.308
26.6085 | -35.5804 | -36.3135
26.6146 | -35.5808 | -36.319
26.6207 | -35.5811 | -36.3245
26.6268 | -35.5815 | -36.3301
26.6329 | -35.5818 | -36.3356
26.639 | -35.5822 | -36.341
26.6451 | -35.5825 | -36.3462
26.6512 | -35.5828 | -36.3513
26.6573 | -35.5831 | -36.3568
26.6634 | -35.5834 | -36.3612
26.6695 | -35.5837 | -36.3665
26.6756 | -35.5841 | -36.372
26.6817 | -35.5845 | -36.3774
26.6878 | -35.5848 | -36.3829
26.6939 | -35.5852 | -36.3882
26.7 | -35.5856 | -36.3933
26.7061 | -35.5859 | -36.3985
26.7122 | -35.5863 | -36.4038
26.7183 | -35.5866 | -36.409
26.7245 | -35.587 | -36.414
26.7306 | -35.5874 | -36.419
26.7367 | -35.5877 | -36.424
26.7428 | -35.588 | -36.4291
26.7489 | -35.5884 | -36.4343
26.755 | -35.5887 | -36.4398
26.7611 | -35.5891 | -36.4452
26.7672 | -35.5895 | -36.4506
26.7733 | -35.5898 | -36.4559
26.7794 | -35.5902 | -36.4611
26.7855 | -35.5906 | -36.4664
26.7916 | -35.591 | -36.4714
26.7977 | -35.5913 | -36.4767
26.8038 | -35.5917 | -36.482
26.8099 | -35.5921 | -36.4873
26.816 | -35.5924 | -36.4924
26.8221 | -35.5927 | -36.4976
26.8282 | -35.5931 | -36.5027
26.8343 | -35.5934 | -36.508
26.8404 | -35.5938 | -36.5132
26.8465 | -35.5941 | -36.5185
26.8526 | -35.5945 | -36.5236
26.8587 | -35.5948 | -36.529
26.8649 | -35.5951 | -36.5341
26.871 | -35.5955 | -36.5395
26.8771 | -35.5958 | -36.5446
26.8832 | -35.5961 | -36.5497
26.8893 | -35.5964 | -36.555
26.8954 | -35.5968 | -36.5603
26.9015 | -35.5971 | -36.566
26.9076 | -35.5975 | -36.5715
26.9137 | -35.5979 | -36.5766
26.9198 | -35.5983 | -36.5821
26.9259 | -35.5987 | -36.5874
26.932 | -35.599 | -36.5929
26.9381 | -35.5993 | -36.5984
26.9442 | -35.5996 | -36.6035
26.9503 | -35.5999 | -36.6088
26.9564 | -35.6003 | -36.6142
26.9625 | -35.6007 | -36.6195
26.9686 | -35.6011 | -36.6248
26.9747 | -35.6014 | -36.6298
26.9808 | -35.6018 | -36.6347
26.9869 | -35.6021 | -36.6398
26.993 | -35.6025 | -36.6447
26.9991 | -35.6029 | -36.6498
27.0052 | -35.6032 | -36.6553
27.0114 | -35.6035 | -36.6608
27.0175 | -35.6039 | -36.666
27.0236 | -35.6042 | -36.671
27.0297 | -35.6045 | -36.676
27.0358 | -35.6049 | -36.6811
27.0419 | -35.6052 | -36.6866
27.048 | -35.6056 | -36.6918
27.0541 | -35.6059 | -36.6972
27.0602 | -35.6063 | -36.7023
27.0663 | -35.6066 | -36.7078
27.0724 | -35.6069 | -36.713
27.0785 | -35.6073 | -36.7183
27.0846 | -35.6076 | -36.7231
27.0907 | -35.608 | -36.7286
27.0968 | -35.6083 | -36.7341
27.1029 | -35.6087 | -36.7393
27.109 | -35.609 | -36.7445
27.1151 | -35.6093 | -36.7494
27.1212 | -35.6096 | -36.7545
27.1273 | -35.61 | -36.7598
27.1334 | -35.6104 | -36.7651
27.1395 | -35.6107 | -36.7702
27.1456 | -35.6111 | -36.7753
27.1518 | -35.6115 | -36.7804
27.1579 | -35.6118 | -36.7853
27.164 | -35.6121 | -36.7904
27.1701 | -35.6125 | -36.7956
27.1762 | -35.6128 | -36.801
27.1823 | -35.6131 | -36.8065
27.1884 | -35.6134 | -36.8116
27.1945 | -35.6138 | -36.8162
27.2006 | -35.6141 | -36.8215
27.2067 | -35.6144 | -36.8267
27.2128 | -35.6149 | -36.8318
27.2189 | -35.6152 | -36.8369
27.225 | -35.6155 | -36.8423
27.2311 | -35.6159 | -36.8472
27.2372 | -35.6162 | -36.8525
27.2433 | -35.6166 | -36.8575
27.2494 | -35.617 | -36.8628
27.2555 | -35.6173 | -36.868
27.2616 | -35.6176 | -36.8731
27.2677 | -35.6179 | -36.8781
27.2738 | -35.6183 | -36.8833
27.2799 | -35.6186 | -36.8885
27.286 | -35.619 | -36.8937
27.2921 | -35.6193 | -36.8988
27.2983 | -35.6197 | -36.9036
27.3044 | -35.62 | -36.9087
27.3105 | -35.6203 | -36.9138
27.3166 | -35.6207 | -36.9189
27.3227 | -35.621 | -36.924
27.3288 | -35.6214 | -36.9292
27.3349 | -35.6217 | -36.9345
27.341 | -35.622 | -36.9398
27.3471 | -35.6224 | -36.9447
27.3532 | -35.6227 | -36.9495
27.3593 | -35.6231 | -36.9551
27.3654 | -35.6234 | -36.9603
27.3715 | -35.6238 | -36.9652
27.3776 | -35.6242 | -36.9705
27.3837 | -35.6245 | -36.9756
27.3898 | -35.6249 | -36.9806
27.3959 | -35.6252 | -36.9859
27.402 | -35.6255 | -36.9908
27.4081 | -35.6259 | -36.996
27.4142 | -35.6262 | -37.0013
27.4203 | -35.6265 | -37.0065
27.4264 | -35.6268 | -37.0115
27.4325 | -35.6272 | -37.0166
27.4387 | -35.6274 | -37.0215
27.4448 | -35.6277 | -37.0267
27.4509 | -35.6281 | -37.0319
27.457 | -35.6285 | -37.0372
27.4631 | -35.6288 | -37.042
27.4692 | -35.6291 | -37.0467
27.4753 | -35.6294 | -37.0515
27.4814 | -35.6298 | -37.0566
27.4875 | -35.6301 | -37.0617
27.4936 | -35.6305 | -37.0666
27.4997 | -35.6308 | -37.0719
27.5058 | -35.6311 | -37.0771
27.5119 | -35.6314 | -37.0825
27.518 | -35.6317 | -37.0879
27.5241 | -35.6321 | -37.093
27.5302 | -35.6324 | -37.0985
27.5363 | -35.6327 | -37.1039
27.5424 | -35.633 | -37.109
27.5485 | -35.6333 | -37.1146
27.5546 | -35.6337 | -37.1197
27.5607 | -35.634 | -37.1247
27.5668 | -35.6343 | -37.1294
27.5729 | -35.6347 | -37.1345
27.5791 | -35.635 | -37.1394
27.5852 | -35.6353 | -37.1446
27.5913 | -35.6357 | -37.1497
27.5974 | -35.6361 | -37.155
27.6035 | -35.6364 | -37.1602
27.6096 | -35.6367 | -37.1652
27.6157 | -35.637 | -37.1702
27.6218 | -35.6373 | -37.1755
27.6279 | -35.6376 | -37.1808
27.634 | -35.638 | -37.186
27.6401 | -35.6383 | -37.1914
27.6462 | -35.6387 | -37.1966
27.6523 | -35.639 | -37.2014
27.6584 | -35.6393 | -37.2064
27.6645 | -35.6397 | -37.2116
27.6706 | -35.64 | -37.2168
27.6767 | -35.6404 | -37.2216
27.6828 | -35.6407 | -37.2264
27.6889 | -35.6411 | -37.2313
27.695 | -35.6414 | -37.2364
27.7011 | -35.6417 | -37.2413
27.7072 | -35.642 | -37.2465
27.7133 | -35.6424 | -37.2516
27.7194 | -35.6427 | -37.2564
27.7256 | -35.643 | -37.2616
27.7317 | -35.6433 | -37.2667
27.7378 | -35.6436 | -37.2718
27.7439 | -35.6439 | -37.2771
27.75 | -35.6443 | -37.2823
27.7561 | -35.6446 | -37.2874
27.7622 | -35.6449 | -37.2924
27.7683 | -35.6452 | -37.2975
27.7744 | -35.6456 | -37.3027
27.7805 | -35.6459 | -37.3078
27.7866 | -35.6462 | -37.313
27.7927 | -35.6466 | -37.318
27.7988 | -35.6469 | -37.323
27.8049 | -35.6472 | -37.3279
27.811 | -35.6475 | -37.3329
27.8171 | -35.6479 | -37.3377
27.8232 | -35.6482 | -37.3426
27.8293 | -35.6486 | -37.3475
27.8354 | -35.6489 | -37.3526
27.8415 | -35.6492 | -37.3576
27.8476 | -35.6495 | -37.3627
27.8537 | -35.6498 | -37.3679
27.8598 | -35.6502 | -37.3727
27.866 | -35.6505 | -37.3776
27.8721 | -35.6508 | -37.3827
27.8782 | -35.6511 | -37.3879
27.8843 | -35.6514 | -37.3931
27.8904 | -35.6518 | -37.3981
27.8965 | -35.6521 | -37.403
27.9026 | -35.6524 | -37.408
27.9087 | -35.6528 | -37.413
27.9148 | -35.653 | -37.4177
27.9209 | -35.6533 | -37.4228
27.927 | -35.6536 | -37.4277
27.9331 | -35.654 | -37.4328
27.9392 | -35.6543 | -37.4379
27.9453 | -35.6546 | -37.4429
27.9514 | -35.6549 | -37.448
27.9575 | -35.6553 | -37.4533
27.9636 | -35.6556 | -37.4584
27.9697 | -35.6559 | -37.463
27.9758 | -35.6562 | -37.4682
27.9819 | -35.6565 | -37.4734
27.988 | -35.6568 | -37.4784
27.9941 | -35.6572 | -37.4836
28.0002 | -35.6575 | -37.4888
28.0063 | -35.6578 | -37.494
28.0125 | -35.6581 | -37.4989
28.0186 | -35.6584 | -37.5043
28.0247 | -35.6588 | -37.5094
28.0308 | -35.6591 | -37.5143
28.0369 | -35.6594 | -37.5193
28.043 | -35.6597 | -37.5245
28.0491 | -35.6601 | -37.5297
28.0552 | -35.6604 | -37.5349
28.0613 | -35.6607 | -37.5401
28.0674 | -35.661 | -37.545
28.0735 | -35.6612 | -37.5499

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.6056 | -41.8639
21.4908 | -35.5894 | -34.4537
16.4948 | -34.966 | -27.1822
12.6602 | -34.604 | -20.9866
9.7171 | -34.3856 | -16.1624
7.4582 | -34.2378 | -12.2379
5.7244 | -34.1312 | -9.0593
4.3936 | -34.0682 | -6.4927
3.3722 | -34.0567 | -4.4711
2.5883 | -34.0401 | -2.9412
1.9866 | -34.0289 | -1.7153
1.5248 | -34.0229 | -0.73198
1.1703 | -34.0267 | -0.095755
1 | -34.0162 | 0.21885
0.89822 | -34.0056 | 0.48701
0.68941 | -33.9905 | 1.0601
0.52919 | -33.9734 | 1.501
0.40612 | -33.999 | 2.1039
0.31171 | -34.0042 | 2.5451
0.23928 | -34.01 | 3.2679
0.18358 | -33.9941 | 3.6991
0.14091 | -34.0334 | 4.578
0.10819 | -34.0589 | 5.7387
0.083017 | -34.0947 | 7.21
0.063705 | -34.1273 | 9.0796
0.048923 | -34.182 | 11.5728
0.037479 | -34.3164 | 14.759
0.028801 | -34.4818 | 18.7736
0.022078 | -34.7925 | 24.1198
0.017023 | -35.2711 | 30.5317
0.013018 | -36.0664 | 38.2077
0.010014 | -37.1844 | 46.096

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003727 | -6.5862
0.007454 | -25.6218
0.011181 | -34.111
0.014908 | -40.6423
0.018635 | -42.2559
0.022362 | -44.5523
0.026089 | -45.8052
0.029816 | -46.8773
0.033543 | -47.2375
0.03727 | -48.4923
0.040997 | -50.7345
0.044724 | -52.7842
0.048451 | -53.7973
0.052178 | -55.5245
0.055905 | -54.6519
0.059632 | -55.2112
0.063359 | -56.9548
0.067086 | -57.628
0.070813 | -58.418
0.07454 | -59.2814
0.078267 | -58.6044
0.081994 | -58.46
0.085721 | -60.0407
0.089448 | -60.8687
0.093175 | -60.784
0.096902 | -61.8242
0.10063 | -62.2509
0.10436 | -62.1994
0.10808 | -62.6659
0.11181 | -63.7386
0.11554 | -64.7699
0.11926 | -64.8659
0.12299 | -65.807
0.12672 | -66.856
0.13045 | -67.1079
0.13417 | -66.3792
0.1379 | -65.7919
0.14163 | -65.5042
0.14535 | -65.9722
0.14908 | -66.7655
0.15281 | -67.3554
0.15653 | -67.998
0.16026 | -67.4501
0.16399 | -67.8122
0.16772 | -68.4908
0.17144 | -69.204
0.17517 | -70.1898
0.1789 | -68.43
0.18262 | -68.9711
0.18635 | -69.928
0.19008 | -70.3336
0.1938 | -70.9487
0.19753 | -70.5567
0.20126 | -69.8846
0.20499 | -70.4199
0.20871 | -71.0849
0.21244 | -71.3709
0.21617 | -70.8358
0.21989 | -71.3703
0.22362 | -71.1661
0.22735 | -71.5053
0.23107 | -72.291
0.2348 | -71.3606
0.23853 | -71.545
0.24225 | -72.2826
0.24598 | -72.2211
0.24971 | -72.8097
0.25344 | -73.1576
0.25716 | -73.2957
0.26089 | -73.8172
0.26462 | -72.9417
0.26834 | -72.6408
0.27207 | -73.4439
0.2758 | -73.5027
0.27953 | -73.7305
0.28325 | -74.2667
0.28698 | -74.0762
0.29071 | -74.0481
0.29443 | -74.4337
0.29816 | -74.5674
0.30189 | -74.1105
0.30561 | -73.6885
0.30934 | -74.0816
0.31307 | -74.5903
0.31679 | -75.0312
0.32052 | -75.5613
0.32425 | -75.9483
0.32798 | -77.1023
0.3317 | -76.0733
0.33543 | -75.4057
0.33916 | -75.4859
0.34288 | -75.8957
0.34661 | -76.144
0.35034 | -77.0256
0.35407 | -76.5633
0.35779 | -75.7775
0.36152 | -76.4417
0.36525 | -76.0554
0.36897 | -75.6476
0.3727 | -75.9813
0.37643 | -77.1893
0.38015 | -77.3902
0.38388 | -77.2896
0.38761 | -77.7104
0.39133 | -77.1869
0.39506 | -76.8186
0.39879 | -76.7255
0.40252 | -76.7198
0.40624 | -76.8036
0.40997 | -78.2039
0.4137 | -78.7058
0.41742 | -78.1285
0.42115 | -78.4585
0.42488 | -77.4081
0.42861 | -78.1947
0.43233 | -78.2398
0.43606 | -78.656
0.43979 | -78.3646
0.44351 | -77.9115
0.44724 | -77.8429
0.45097 | -78.3807
0.45469 | -78.2619
0.45842 | -78.3285
0.46215 | -78.1814
0.46587 | -78.0516
0.4696 | -78.2741
0.47333 | -78.7558
0.47706 | -79.2477
0.48078 | -79.5264
0.48451 | -79.9561
0.48824 | -80.7112
0.49196 | -79.594
0.49569 | -80.8968
0.49942 | -80.9285
0.50314 | -80.1264
0.50687 | -79.5474
0.5106 | -79.8017
0.51433 | -80.0757
0.51805 | -81.2611
0.52178 | -81.7065
0.52551 | -80.9886
0.52923 | -80.8943
0.53296 | -81.2864
0.53669 | -80.979
0.54041 | -80.8107
0.54414 | -80.7738
0.54787 | -81.0009
0.5516 | -81.2582
0.55532 | -80.9444
0.55905 | -80.3363
0.56278 | -80.1207
0.5665 | -80.2541
0.57023 | -80.4942
0.57396 | -80.6825
0.57769 | -80.6877
0.58141 | -80.4403
0.58514 | -80.2563
0.58887 | -80.4499
0.59259 | -80.472
0.59632 | -80.2341
0.60005 | -80.5304
0.60377 | -81.331
0.6075 | -82.0176
0.61123 | -82.096
0.61496 | -81.5224
0.61868 | -81.1972
0.62241 | -81.5691
0.62614 | -81.7117
0.62986 | -81.6066
0.63359 | -81.884
0.63732 | -82.0089
0.64104 | -82.0593
0.64477 | -82.4098
0.6485 | -82.4457
0.65223 | -82.3506
0.65595 | -82.5189
0.65968 | -82.7381
0.66341 | -82.8301
0.66713 | -82.9915
0.67086 | -82.9978
0.67459 | -82.5336
0.67831 | -82.1061
0.68204 | -82.1786
0.68577 | -82.1299
0.68949 | -81.8188
0.69322 | -81.8876
0.69695 | -82.3485
0.70068 | -82.7785
0.7044 | -83.1678
0.70813 | -83.5229
0.71186 | -83.5692
0.71558 | -83.473
0.71931 | -83.7292
0.72304 | -83.6456
0.72676 | -82.9912
0.73049 | -82.0941
0.73422 | -81.585
0.73794 | -81.8986
0.74167 | -82.8606
0.7454 | -83.7002
0.74913 | -83.9834
0.75285 | -83.7387
0.75658 | -83.3924
0.76031 | -83.1716
0.76403 | -83.1528
0.76776 | -83.2163
0.77149 | -83.4608
0.77521 | -84.0641
0.77894 | -84.1435
0.78267 | -83.4505
0.7864 | -83.205
0.79012 | -83.5597
0.79385 | -83.7169
0.79758 | -83.7037
0.8013 | -83.889
0.80503 | -84.0556
0.80876 | -84.0322
0.81249 | -83.915
0.81621 | -84.0268
0.81994 | -84.299
0.82367 | -84.6093
0.82739 | -84.6434
0.83112 | -84.3828
0.83485 | -84.4343
0.83857 | -84.9569
0.8423 | -85.1975
0.84603 | -84.99
0.84976 | -84.7898
0.85348 | -85.0428
0.85721 | -85.9661
0.86094 | -86.217
0.86466 | -85.3444
0.86839 | -84.6505
0.87212 | -84.4284
0.87584 | -84.3735
0.87957 | -84.1657
0.8833 | -83.9248
0.88702 | -84.0716
0.89075 | -84.5435
0.89448 | -84.7753
0.89821 | -84.5554
0.90193 | -84.7373
0.90566 | -85.2372
0.90939 | -85.2455
0.91311 | -84.9438
0.91684 | -84.8296
0.92057 | -85.2019
0.92429 | -85.5738
0.92802 | -85.5623
0.93175 | -85.5103
0.93548 | -85.8255
0.9392 | -85.8578
0.94293 | -85.5861
0.94666 | -85.3832
0.95038 | -85.5531
0.95411 | -85.9263
0.95784 | -86.0306
0.96157 | -85.9065
0.96529 | -85.6692
0.96902 | -85.3469
0.97275 | -85.1019
0.97647 | -85.2901
0.9802 | -85.5572
0.98393 | -85.6477
0.98765 | -85.6544
0.99138 | -86.0238
0.99511 | -86.2889
0.99884 | -86.0958
1.0026 | -85.9252
1.0063 | -86.1383
1.01 | -86.2816
1.0137 | -86.1983
1.0175 | -85.791
1.0212 | -85.4295
1.0249 | -85.3559
1.0287 | -85.4557
1.0324 | -85.8242
1.0361 | -86.2444
1.0398 | -86.6169
1.0436 | -86.7665
1.0473 | -86.5969
1.051 | -86.3226
1.0547 | -86.1693
1.0585 | -86.0309
1.0622 | -85.9306
1.0659 | -86.1098
1.0696 | -86.4745
1.0734 | -86.5087
1.0771 | -86.5128
1.0808 | -86.4756
1.0846 | -86.4617
1.0883 | -86.5407
1.092 | -86.7033
1.0957 | -86.7296
1.0995 | -86.4846
1.1032 | -86.2247
1.1069 | -86.2269
1.1106 | -86.4235
1.1144 | -86.5636
1.1181 | -86.5727
1.1218 | -86.5728
1.1256 | -86.5578
1.1293 | -86.3927
1.133 | -86.1151
1.1367 | -85.9883
1.1405 | -86.2235
1.1442 | -86.7031
1.1479 | -87.0233
1.1516 | -87.0265
1.1554 | -87.0114
1.1591 | -87.2021
1.1628 | -87.4679
1.1665 | -87.5032
1.1703 | -87.2198
1.174 | -86.8071
1.1777 | -86.5227
1.1815 | -86.5189
1.1852 | -86.7868
1.1889 | -87.1828
1.1926 | -87.5808
1.1964 | -87.9383
1.2001 | -88.1442
1.2038 | -88.0649
1.2075 | -87.7529
1.2113 | -87.4872
1.215 | -87.5426
1.2187 | -87.9072
1.2225 | -88.3486
1.2262 | -88.6413
1.2299 | -88.7027
1.2336 | -88.6238
1.2374 | -88.5145
1.2411 | -88.426
1.2448 | -88.3717
1.2485 | -88.3554
1.2523 | -88.392
1.256 | -88.4848
1.2597 | -88.5034
1.2635 | -88.2867
1.2672 | -87.9438
1.2709 | -87.7204
1.2746 | -87.7286
1.2784 | -87.9445
1.2821 | -88.2972
1.2858 | -88.7274
1.2895 | -89.0878
1.2933 | -89.1814
1.297 | -89.0679
1.3007 | -88.9914
1.3044 | -88.9931
1.3082 | -88.8512
1.3119 | -88.5612
1.3156 | -88.4001
1.3194 | -88.4024
1.3231 | -88.4682
1.3268 | -88.5795
1.3305 | -88.6791
1.3343 | -88.7298
1.338 | -88.7386
1.3417 | -88.6548
1.3454 | -88.4748
1.3492 | -88.3714
1.3529 | -88.4192
1.3566 | -88.4564
1.3604 | -88.5052
1.3641 | -88.731
1.3678 | -89.0279
1.3715 | -89.0917
1.3753 | -88.7919
1.379 | -88.3058
1.3827 | -87.921
1.3864 | -87.8093
1.3902 | -88.0819
1.3939 | -88.7523
1.3976 | -89.4458
1.4014 | -89.5065
1.4051 | -88.8972
1.4088 | -88.228
1.4125 | -87.9458
1.4163 | -88.1499
1.42 | -88.7197
1.4237 | -89.3547
1.4274 | -89.69
1.4312 | -89.5918
1.4349 | -89.1611
1.4386 | -88.6053
1.4423 | -88.2226
1.4461 | -88.183
1.4498 | -88.4376
1.4535 | -88.8748
1.4573 | -89.3357
1.461 | -89.6234
1.4647 | -89.6606
1.4684 | -89.5746
1.4722 | -89.5365
1.4759 | -89.6111
1.4796 | -89.8018
1.4833 | -90.1076
1.4871 | -90.4403
1.4908 | -90.5852
1.4945 | -90.4086
1.4983 | -90.099
1.502 | -89.9337
1.5057 | -89.9598
1.5094 | -89.9764
1.5132 | -89.7933
1.5169 | -89.3898
1.5206 | -88.9328
1.5243 | -88.6828
1.5281 | -88.7976
1.5318 | -89.1704
1.5355 | -89.4797
1.5393 | -89.6179
1.543 | -89.7485
1.5467 | -89.8904
1.5504 | -89.9256
1.5542 | -89.8605
1.5579 | -89.85
1.5616 | -89.9935
1.5653 | -90.2116
1.5691 | -90.328
1.5728 | -90.3
1.5765 | -90.2397
1.5802 | -90.1538
1.584 | -90.0672
1.5877 | -90.166
1.5914 | -90.4721
1.5952 | -90.6882
1.5989 | -90.5803
1.6026 | -90.2438
1.6063 | -89.9033
1.6101 | -89.7976
1.6138 | -89.7574
1.6175 | -89.8481
1.6212 | -90.0194
1.625 | -90.2125
1.6287 | -90.3863
1.6324 | -90.5304
1.6362 | -90.6408
1.6399 | -90.717
1.6436 | -90.7986
1.6473 | -90.8987
1.6511 | -90.933
1.6548 | -90.8176
1.6585 | -90.5613
1.6622 | -90.3225
1.666 | -90.2742
1.6697 | -90.4484
1.6734 | -90.7294
1.6771 | -90.9458
1.6809 | -91.0506
1.6846 | -91.1094
1.6883 | -91.1229
1.6921 | -91.0579
1.6958 | -90.9126
1.6995 | -90.7296
1.7032 | -90.5813
1.707 | -90.5539
1.7107 | -90.6882
1.7144 | -90.93
1.7181 | -91.1834
1.7219 | -91.3732
1.7256 | -91.427
1.7293 | -91.3056
1.7331 | -91.0602
1.7368 | -90.8094
1.7405 | -90.6447
1.7442 | -90.5841
1.748 | -90.5824
1.7517 | -90.5814
1.7554 | -90.5709
1.7591 | -90.5759
1.7629 | -90.6233
1.7666 | -90.7065
1.7703 | -90.7818
1.7741 | -90.7848
1.7778 | -90.6612
1.7815 | -90.4512
1.7852 | -90.312
1.789 | -90.3667
1.7927 | -90.6007
1.7964 | -90.8994
1.8001 | -91.1557
1.8039 | -91.3294
1.8076 | -91.4274
1.8113 | -91.4679
1.815 | -91.4729
1.8188 | -91.4233
1.8225 | -91.33
1.8262 | -91.2681
1.83 | -91.3173
1.8337 | -91.4637
1.8374 | -91.5907
1.8411 | -91.6093
1.8449 | -91.5502
1.8486 | -91.4655
1.8523 | -91.3858
1.856 | -91.3166
1.8598 | -91.2517
1.8635 | -91.1797
1.8672 | -91.1468
1.871 | -91.2126
1.8747 | -91.3317
1.8784 | -91.3597
1.8821 | -91.2484
1.8859 | -91.0835
1.8896 | -90.9847
1.8933 | -90.9975
1.897 | -91.0525
1.9008 | -91.0342
1.9045 | -90.9135
1.9082 | -90.7657
1.912 | -90.6605
1.9157 | -90.5982
1.9194 | -90.5836
1.9231 | -90.667
1.9269 | -90.9159
1.9306 | -91.3457
1.9343 | -91.8782
1.938 | -92.3645
1.9418 | -92.624
1.9455 | -92.5579
1.9492 | -92.2899
1.9529 | -92.0266
1.9567 | -91.8663
1.9604 | -91.8005
1.9641 | -91.7897
1.9679 | -91.8263
1.9716 | -91.8952
1.9753 | -91.9616
1.979 | -91.9504
1.9828 | -91.8053
1.9865 | -91.5872
1.9902 | -91.4391
1.9939 | -91.4291
1.9977 | -91.5029
2.0014 | -91.5537
2.0051 | -91.533
2.0089 | -91.4746
2.0126 | -91.4455
2.0163 | -91.4746
2.02 | -91.5482
2.0238 | -91.6429
2.0275 | -91.7345
2.0312 | -91.7875
2.0349 | -91.7557
2.0387 | -91.6419
2.0424 | -91.5259
2.0461 | -91.4901
2.0498 | -91.5195
2.0536 | -91.5407
2.0573 | -91.554
2.061 | -91.6255
2.0648 | -91.7922
2.0685 | -92.0428
2.0722 | -92.3126
2.0759 | -92.5194
2.0797 | -92.6336
2.0834 | -92.6915
2.0871 | -92.7155
2.0908 | -92.6658
2.0946 | -92.5274
2.0983 | -92.3443
2.102 | -92.2118
2.1058 | -92.1683
2.1095 | -92.1443
2.1132 | -92.0932
2.1169 | -92.0698
2.1207 | -92.1203
2.1244 | -92.1998
2.1281 | -92.2162
2.1318 | -92.176
2.1356 | -92.1791
2.1393 | -92.2801
2.143 | -92.4043
2.1468 | -92.4442
2.1505 | -92.3836
2.1542 | -92.3688
2.1579 | -92.3656
2.1617 | -92.3586
2.1654 | -92.3248
2.1691 | -92.2562
2.1728 | -92.1707
2.1766 | -92.1118
2.1803 | -92.1173
2.184 | -92.2068
2.1877 | -92.3945
2.1915 | -92.6439
2.1952 | -92.8864
2.1989 | -93.0568
2.2027 | -93.1354
2.2064 | -93.1491
2.2101 | -93.1322
2.2138 | -93.1044
2.2176 | -93.0674
2.2213 | -93.0182
2.225 | -92.9349
2.2287 | -92.8042
2.2325 | -92.6633
2.2362 | -92.5783
2.2399 | -92.5975
2.2437 | -92.7271
2.2474 | -92.9271
2.2511 | -93.1278
2.2548 | -93.2733
2.2586 | -93.3607
2.2623 | -93.4133
2.266 | -93.4556
2.2697 | -93.4633
2.2735 | -93.3986
2.2772 | -93.2759
2.2809 | -93.1454
2.2846 | -93.0406
2.2884 | -92.968
2.2921 | -92.9219
2.2958 | -92.8907
2.2996 | -92.8858
2.3033 | -92.9415
2.307 | -93.0822
2.3107 | -93.3099
2.3145 | -93.6009
2.3182 | -93.8682
2.3219 | -94.0313
2.3256 | -94.0607
2.3294 | -93.9904
2.3331 | -93.8787
2.3368 | -93.7736
2.3406 | -93.6612
2.3443 | -93.5228
2.348 | -93.3772
2.3517 | -93.2848
2.3555 | -93.3048
2.3592 | -93.4542
2.3629 | -93.676
2.3666 | -93.875
2.3704 | -93.9586
2.3741 | -93.8863
2.3778 | -93.6885
2.3816 | -93.4377
2.3853 | -93.218
2.389 | -93.0919
2.3927 | -93.0624
2.3965 | -93.0923
2.4002 | -93.1303
2.4039 | -93.1182
2.4076 | -93.0429
2.4114 | -92.9489
2.4151 | -92.8979
2.4188 | -92.9273
2.4225 | -93.0177
2.4263 | -93.1193
2.43 | -93.2036
2.4337 | -93.2707
2.4375 | -93.3181
2.4412 | -93.3387
2.4449 | -93.3448
2.4486 | -93.3654
2.4524 | -93.4241
2.4561 | -93.5133
2.4598 | -93.603
2.4635 | -93.6798
2.4673 | -93.7425
2.471 | -93.8029
2.4747 | -93.8764
2.4785 | -93.9589
2.4822 | -94.0386
2.4859 | -94.1073
2.4896 | -94.1423
2.4934 | -94.1245
2.4971 | -94.0647
2.5008 | -93.997
2.5045 | -93.9483
2.5083 | -93.8956
2.512 | -93.8058
2.5157 | -93.6982
2.5194 | -93.6265
2.5232 | -93.6239
2.5269 | -93.6506
2.5306 | -93.675
2.5344 | -93.6824
2.5381 | -93.6742
2.5418 | -93.666
2.5455 | -93.6722
2.5493 | -93.692
2.553 | -93.7172
2.5567 | -93.7196
2.5604 | -93.6935
2.5642 | -93.6606
2.5679 | -93.6598
2.5716 | -93.6872
2.5754 | -93.7351
2.5791 | -93.7759
2.5828 | -93.7826
2.5865 | -93.743
2.5903 | -93.6644
2.594 | -93.5737
2.5977 | -93.5131
2.6014 | -93.5281
2.6052 | -93.6317
2.6089 | -93.7999
2.6126 | -93.9835
2.6164 | -94.1174
2.6201 | -94.1759
2.6238 | -94.1797
2.6275 | -94.179
2.6313 | -94.2204
2.635 | -94.3177
2.6387 | -94.4643
2.6424 | -94.6203
2.6462 | -94.773
2.6499 | -94.9225
2.6536 | -95.0513
2.6573 | -95.1376
2.6611 | -95.1267
2.6648 | -95.0036
2.6685 | -94.8077
2.6723 | -94.5976
2.676 | -94.3975
2.6797 | -94.2278
2.6834 | -94.0967
2.6872 | -94.1299
2.6909 | -94.1706
2.6946 | -94.281
2.6983 | -94.4239
2.7021 | -94.5579
2.7058 | -94.6622
2.7095 | -94.7322
2.7133 | -94.7736
2.717 | -94.7944
2.7207 | -94.8027
2.7244 | -94.8102
2.7282 | -94.8322
2.7319 | -94.8385
2.7356 | -94.8157
2.7393 | -94.7534
2.7431 | -94.6691
2.7468 | -94.59
2.7505 | -94.5405
2.7543 | -94.5274
2.758 | -94.5322
2.7617 | -94.5364
2.7654 | -94.5291
2.7692 | -94.5342
2.7729 | -94.5636
2.7766 | -94.6086
2.7803 | -94.6373
2.7841 | -94.6156
2.7878 | -94.5324
2.7915 | -94.4014
2.7952 | -94.2629
2.799 | -94.1487
2.8027 | -94.0985
2.8064 | -94.1335
2.8102 | -94.2447
2.8139 | -94.3809
2.8176 | -94.4802
2.8213 | -94.5117
2.8251 | -94.4822
2.8288 | -94.4077
2.8325 | -94.3129
2.8362 | -94.2269
2.84 | -94.1838
2.8437 | -94.2076
2.8474 | -94.31
2.8512 | -94.4848
2.8549 | -94.6808
2.8586 | -94.8437
2.8623 | -94.9401
2.8661 | -94.9541
2.8698 | -94.9009
2.8735 | -94.8009
2.8772 | -94.6827
2.881 | -94.5569
2.8847 | -94.428
2.8884 | -94.3058
2.8921 | -94.2134
2.8959 | -94.1814
2.8996 | -94.2146
2.9033 | -94.3165
2.9071 | -94.4475
2.9108 | -94.586
2.9145 | -94.7212
2.9182 | -94.8329
2.922 | -94.9015
2.9257 | -94.9066
2.9294 | -94.8546
2.9331 | -94.8006
2.9369 | -94.8027
2.9406 | -94.8851
2.9443 | -95.0368
2.9481 | -95.2267
2.9518 | -95.4013
2.9555 | -95.5125
2.9592 | -95.5075
2.963 | -95.4045
2.9667 | -95.2654
2.9704 | -95.151
2.9741 | -95.102
2.9779 | -95.125
2.9816 | -95.1876
2.9853 | -95.2274
2.9891 | -95.2265
2.9928 | -95.1664
2.9965 | -95.0455
3.0002 | -94.8913
3.004 | -94.7526
3.0077 | -94.6663
3.0114 | -94.6455
3.0151 | -94.688
3.0189 | -94.7826
3.0226 | -94.9046
3.0263 | -95.034
3.03 | -95.1567
3.0338 | -95.2803
3.0375 | -95.3895
3.0412 | -95.4871
3.045 | -95.5717
3.0487 | -95.6454
3.0524 | -95.7038
3.0561 | -95.7516
3.0599 | -95.7955
3.0636 | -95.8259
3.0673 | -95.832
3.071 | -95.8271
3.0748 | -95.8201
3.0785 | -95.7942
3.0822 | -95.7339
3.086 | -95.6514
3.0897 | -95.5723
3.0934 | -95.5084
3.0971 | -95.4578
3.1009 | -95.4179
3.1046 | -95.3588
3.1083 | -95.2513
3.112 | -95.113
3.1158 | -94.9833
3.1195 | -94.8929
3.1232 | -94.861
3.127 | -94.9022
3.1307 | -94.9927
3.1344 | -95.097
3.1381 | -95.2021
3.1419 | -95.311
3.1456 | -95.4146
3.1493 | -95.5047
3.153 | -95.5805
3.1568 | -95.6377
3.1605 | -95.6665
3.1642 | -95.6692
3.1679 | -95.6528
3.1717 | -95.6181
3.1754 | -95.5564
3.1791 | -95.4836
3.1829 | -95.4174
3.1866 | -95.3651
3.1903 | -95.3355
3.194 | -95.3238
3.1978 | -95.3326
3.2015 | -95.3663
3.2052 | -95.4282
3.2089 | -95.5121
3.2127 | -95.6046
3.2164 | -95.6946
3.2201 | -95.7708
3.2239 | -95.8101
3.2276 | -95.8648
3.2313 | -95.9169
3.235 | -95.9702
3.2388 | -96.0331
3.2425 | -96.1035
3.2462 | -96.1923
3.2499 | -96.2859
3.2537 | -96.3772
3.2574 | -96.4493
3.2611 | -96.5101
3.2649 | -96.556
3.2686 | -96.5826
3.2723 | -96.5936
3.276 | -96.575
3.2798 | -96.5292
3.2835 | -96.441
3.2872 | -96.3276
3.2909 | -96.2067
3.2947 | -96.0806
3.2984 | -95.9716
3.3021 | -95.8804
3.3058 | -95.8203
3.3096 | -95.795
3.3133 | -95.7985
3.317 | -95.8093
3.3208 | -95.8312
3.3245 | -95.8549
3.3282 | -95.867
3.3319 | -95.8725
3.3357 | -95.8713
3.3394 | -95.8883
3.3431 | -95.9316
3.3468 | -95.9993
3.3506 | -96.0956
3.3543 | -96.1974
3.358 | -96.2791
3.3618 | -96.3158
3.3655 | -96.2999
3.3692 | -96.2442
3.3729 | -96.1593
3.3767 | -96.0789
3.3804 | -96.0401
3.3841 | -96.0633
3.3878 | -96.1397
3.3916 | -96.2496
3.3953 | -96.3632
3.399 | -96.4662
3.4027 | -96.5446
3.4065 | -96.5781
3.4102 | -96.5719
3.4139 | -96.5554
3.4177 | -96.5381
3.4214 | -96.5221
3.4251 | -96.5007
3.4288 | -96.4697
3.4326 | -96.4177
3.4363 | -96.3532
3.44 | -96.2905
3.4437 | -96.245
3.4475 | -96.2264
3.4512 | -96.2347
3.4549 | -96.2787
3.4587 | -96.3685
3.4624 | -96.4859
3.4661 | -96.5953
3.4698 | -96.6515
3.4736 | -96.6313
3.4773 | -96.5472
3.481 | -96.425
3.4847 | -96.2993
3.4885 | -96.1802
3.4922 | -96.0959
3.4959 | -96.0599
3.4997 | -96.0698
3.5034 | -96.1195
3.5071 | -96.1928
3.5108 | -96.2691
3.5146 | -96.3428
3.5183 | -96.4131
3.522 | -96.4735
3.5257 | -96.5309
3.5295 | -96.59
3.5332 | -96.6443
3.5369 | -96.6947
3.5406 | -96.7337
3.5444 | -96.7561
3.5481 | -96.7274
3.5518 | -96.6517
3.5556 | -96.5559
3.5593 | -96.4824
3.563 | -96.4566
3.5667 | -96.4751
3.5705 | -96.5294
3.5742 | -96.6021
3.5779 | -96.6609
3.5816 | -96.6783
3.5854 | -96.66
3.5891 | -96.6277
3.5928 | -96.5969
3.5966 | -96.5713
3.6003 | -96.5474
3.604 | -96.5253
3.6077 | -96.4996
3.6115 | -96.4698
3.6152 | -96.4426
3.6189 | -96.4416
3.6226 | -96.4737
3.6264 | -96.5324
3.6301 | -96.6089
3.6338 | -96.6969
3.6375 | -96.7872
3.6413 | -96.8743
3.645 | -96.9617
3.6487 | -97.0589
3.6525 | -97.1598
3.6562 | -97.2461
3.6599 | -97.2972
3.6636 | -97.3018
3.6674 | -97.2585
3.6711 | -97.1661
3.6748 | -97.0359
3.6785 | -96.8881
3.6823 | -96.7498
3.686 | -96.642
3.6897 | -96.5614
3.6935 | -96.5024
3.6972 | -96.4596
3.7009 | -96.431
3.7046 | -96.4252
3.7084 | -96.4515
3.7121 | -96.5038
3.7158 | -96.5554
3.7195 | -96.593
3.7233 | -96.6154
3.727 | -96.6255
3.7307 | -96.635
3.7345 | -96.649
3.7382 | -96.6706
3.7419 | -96.7095
3.7456 | -96.7748
3.7494 | -96.8633
3.7531 | -96.9702
3.7568 | -97.0772
3.7605 | -97.1135
3.7643 | -97.1576
3.768 | -97.1792
3.7717 | -97.1843
3.7754 | -97.1844
3.7792 | -97.1806
3.7829 | -97.1833
3.7866 | -97.196
3.7904 | -97.21
3.7941 | -97.2175
3.7978 | -97.2078
3.8015 | -97.1811
3.8053 | -97.1546
3.809 | -97.1357
3.8127 | -97.1312
3.8164 | -97.1489
3.8202 | -97.1852
3.8239 | -97.222
3.8276 | -97.2478
3.8314 | -97.248
3.8351 | -97.2243
3.8388 | -97.1849
3.8425 | -97.1453
3.8463 | -97.1083
3.85 | -97.0747
3.8537 | -97.0518
3.8574 | -97.0351
3.8612 | -97.0198
3.8649 | -97.0001
3.8686 | -96.9805
3.8723 | -96.9679
3.8761 | -96.9708
3.8798 | -96.9977
3.8835 | -97.0534
3.8873 | -97.1307
3.891 | -97.2249
3.8947 | -97.3283
3.8984 | -97.4296
3.9022 | -97.5125
3.9059 | -97.5651
3.9096 | -97.5891
3.9133 | -97.5792
3.9171 | -97.5268
3.9208 | -97.4566
3.9245 | -97.3685
3.9283 | -97.2683
3.932 | -97.1587
3.9357 | -97.0577
3.9394 | -96.971
3.9432 | -96.9057
3.9469 | -96.8591
3.9506 | -96.8363
3.9543 | -96.8355
3.9581 | -96.8622
3.9618 | -96.9164
3.9655 | -96.9979
3.9693 | -97.1015
3.973 | -97.2307
3.9767 | -97.3783
3.9804 | -97.533
3.9842 | -97.6592
3.9879 | -97.7524
3.9916 | -97.7934
3.9953 | -97.7767
3.9991 | -97.7078
4.0028 | -97.6009
4.0065 | -97.485
4.0102 | -97.378
4.014 | -97.2919
4.0177 | -97.2286
4.0214 | -97.1789
4.0252 | -97.1461
4.0289 | -97.1283
4.0326 | -97.11
4.0363 | -97.0856
4.0401 | -97.0564
4.0438 | -97.0186
4.0475 | -96.9837
4.0512 | -96.9709
4.055 | -96.9869
4.0587 | -97.0305
4.0624 | -97.1039
4.0662 | -97.2111
4.0699 | -97.3306
4.0736 | -97.4601
4.0773 | -97.5912
4.0811 | -97.7156
4.0848 | -97.8351
4.0885 | -97.9297
4.0922 | -98.0056
4.096 | -98.0777
4.0997 | -98.1345
4.1034 | -98.175
4.1072 | -98.1916
4.1109 | -98.1868
4.1146 | -98.1479
4.1183 | -98.0864
4.1221 | -98.0063
4.1258 | -97.9074
4.1295 | -97.7999
4.1332 | -97.7008
4.137 | -97.6091
4.1407 | -97.5298
4.1444 | -97.4758
4.1481 | -97.4415
4.1519 | -97.4302
4.1556 | -97.4272
4.1593 | -97.4309
4.1631 | -97.4468
4.1668 | -97.4804
4.1705 | -97.5317
4.1742 | -97.582
4.178 | -97.6352
4.1817 | -97.6836
4.1854 | -97.7174
4.1891 | -97.7336
4.1929 | -97.7268
4.1966 | -97.7087
4.2003 | -97.6805
4.2041 | -97.645
4.2078 | -97.6109
4.2115 | -97.587
4.2152 | -97.579
4.219 | -97.5906
4.2227 | -97.6225
4.2264 | -97.6725
4.2301 | -97.7325
4.2339 | -97.8006
4.2376 | -97.8769
4.2413 | -97.942
4.245 | -97.9949
4.2488 | -98.0352
4.2525 | -98.0601
4.2562 | -98.0738
4.26 | -98.0843
4.2637 | -98.1083
4.2674 | -98.1415
4.2711 | -98.1881
4.2749 | -98.249
4.2786 | -98.3166
4.2823 | -98.3877
4.286 | -98.4541
4.2898 | -98.5124
4.2935 | -98.551
4.2972 | -98.5656
4.301 | -98.5317
4.3047 | -98.5072
4.3084 | -98.4688
4.3121 | -98.4108
4.3159 | -98.339
4.3196 | -98.2524
4.3233 | -98.1692
4.327 | -98.0872
4.3308 | -98.0148
4.3345 | -97.973
4.3382 | -97.9508
4.342 | -97.9439
4.3457 | -97.9524
4.3494 | -97.973
4.3531 | -97.9912
4.3569 | -98.0016
4.3606 | -98.0052
4.3643 | -97.993
4.368 | -97.9641
4.3718 | -97.9255
4.3755 | -97.8814
4.3792 | -97.8457
4.3829 | -97.8344
4.3867 | -97.8509
4.3904 | -97.8966
4.3941 | -97.9609
4.3979 | -98.0375
4.4016 | -98.118
4.4053 | -98.1898
4.409 | -98.2613
4.4128 | -98.3297
4.4165 | -98.3889
4.4202 | -98.4355
4.4239 | -98.4827
4.4277 | -98.5256
4.4314 | -98.5563
4.4351 | -98.5806
4.4389 | -98.6011
4.4426 | -98.6208
4.4463 | -98.6471
4.45 | -98.6666
4.4538 | -98.677
4.4575 | -98.6798
4.4612 | -98.6748
4.4649 | -98.6619
4.4687 | -98.6471
4.4724 | -98.6289
4.4761 | -98.6043
4.4799 | -98.5647
4.4836 | -98.5062
4.4873 | -98.4368
4.491 | -98.3585
4.4948 | -98.2668
4.4985 | -98.1737
4.5022 | -98.0954
4.5059 | -98.0406
4.5097 | -98.0155
4.5134 | -98.0221
4.5171 | -98.0628
4.5208 | -98.1416
4.5246 | -98.2483
4.5283 | -98.3742
4.532 | -98.5083
4.5358 | -98.6336
4.5395 | -98.7411
4.5432 | -98.8299
4.5469 | -98.9037
4.5507 | -98.9729
4.5544 | -99.0408
4.5581 | -99.0974
4.5618 | -99.1537
4.5656 | -99.2214
4.5693 | -99.2753
4.573 | -99.3142
4.5768 | -99.3343
4.5805 | -99.3317
4.5842 | -99.3113
4.5879 | -99.2689
4.5917 | -99.218
4.5954 | -99.1658
4.5991 | -99.126
4.6028 | -99.0991
4.6066 | -99.0983
4.6103 | -99.1122
4.614 | -99.1291
4.6178 | -99.1427
4.6215 | -99.1353
4.6252 | -99.102
4.6289 | -99.0463
4.6327 | -98.9601
4.6364 | -98.8564
4.6401 | -98.7599
4.6438 | -98.67
4.6476 | -98.5956
4.6513 | -98.54
4.655 | -98.5091
4.6587 | -98.5017
4.6625 | -98.5069
4.6662 | -98.5226
4.6699 | -98.5417
4.6737 | -98.5607
4.6774 | -98.5704
4.6811 | -98.5825
4.6848 | -98.6007
4.6886 | -98.6231
4.6923 | -98.6631
4.696 | -98.7216
4.6997 | -98.7932
4.7035 | -98.8883
4.7072 | -98.9934
4.7109 | -99.1009
4.7147 | -99.2059
4.7184 | -99.3005
4.7221 | -99.375
4.7258 | -99.4206
4.7296 | -99.4354
4.7333 | -99.4285
4.737 | -99.3984
4.7407 | -99.3526
4.7445 | -99.2988
4.7482 | -99.2332
4.7519 | -99.1626
4.7556 | -99.0963
4.7594 | -99.0391
4.7631 | -98.9904
4.7668 | -98.9397
4.7706 | -98.8927
4.7743 | -98.847
4.778 | -98.8068
4.7817 | -98.7833
4.7855 | -98.786
4.7892 | -98.8149
4.7929 | -98.8711
4.7966 | -98.9482
4.8004 | -99.0457
4.8041 | -99.1475
4.8078 | -99.2478
4.8116 | -99.3303
4.8153 | -99.3906
4.819 | -99.4248
4.8227 | -99.4279
4.8265 | -99.4083
4.8302 | -99.3776
4.8339 | -99.3398
4.8376 | -99.3324
4.8414 | -99.3241
4.8451 | -99.3269
4.8488 | -99.3325
4.8526 | -99.3371
4.8563 | -99.3363
4.86 | -99.3286
4.8637 | -99.3111
4.8675 | -99.2866
4.8712 | -99.2556
4.8749 | -99.219
4.8786 | -99.1891
4.8824 | -99.1677
4.8861 | -99.1589
4.8898 | -99.1486
4.8935 | -99.1416
4.8973 | -99.1373
4.901 | -99.1285
4.9047 | -99.1159
4.9085 | -99.0933
4.9122 | -99.0648
4.9159 | -99.0348
4.9196 | -99.0039
4.9234 | -98.9789
4.9271 | -98.9665
4.9308 | -98.968
4.9345 | -98.9706
4.9383 | -98.9682
4.942 | -98.9634
4.9457 | -98.9542
4.9495 | -98.9354
4.9532 | -98.9103
4.9569 | -98.8755
4.9606 | -98.8478
4.9644 | -98.8254
4.9681 | -98.8138
4.9718 | -98.8102
4.9755 | -98.8155
4.9793 | -98.8301
4.983 | -98.8613
4.9867 | -98.9103
4.9905 | -98.9613
4.9942 | -99.0173
4.9979 | -99.076
5.0016 | -99.1324
5.0054 | -99.1815
5.0091 | -99.2182
5.0128 | -99.2378
5.0165 | -99.241
5.0203 | -99.2376
5.024 | -99.2179
5.0277 | -99.1954
5.0314 | -99.1751
5.0352 | -99.1616
5.0389 | -99.1577
5.0426 | -99.1691
5.0464 | -99.192
5.0501 | -99.2287
5.0538 | -99.2836
5.0575 | -99.3552
5.0613 | -99.4349
5.065 | -99.5241
5.0687 | -99.6249
5.0724 | -99.7259
5.0762 | -99.8233
5.0799 | -99.9088
5.0836 | -99.9781
5.0874 | -100.039
5.0911 | -100.0909
5.0948 | -100.1158
5.0985 | -100.1083
5.1023 | -100.0816
5.106 | -100.0326
5.1097 | -99.9639
5.1134 | -99.8845
5.1172 | -99.7901
5.1209 | -99.6997
5.1246 | -99.6275
5.1283 | -99.5636
5.1321 | -99.5143
5.1358 | -99.4864
5.1395 | -99.475
5.1433 | -99.4747
5.147 | -99.4767
5.1507 | -99.4776
5.1544 | -99.4801
5.1582 | -99.4884
5.1619 | -99.5004
5.1656 | -99.5021
5.1693 | -99.5073
5.1731 | -99.5212
5.1768 | -99.5387
5.1805 | -99.5602
5.1843 | -99.5844
5.188 | -99.6102
5.1917 | -99.6295
5.1954 | -99.6507
5.1992 | -99.6704
5.2029 | -99.6886
5.2066 | -99.7087
5.2103 | -99.7337
5.2141 | -99.7665
5.2178 | -99.8041
5.2215 | -99.8437
5.2252 | -99.8894
5.229 | -99.9277
5.2327 | -99.95
5.2364 | -99.9624
5.2402 | -99.969
5.2439 | -99.9671
5.2476 | -99.9523
5.2513 | -99.9316
5.2551 | -99.9068
5.2588 | -99.8697
5.2625 | -99.8269
5.2662 | -99.7933
5.27 | -99.7687
5.2737 | -99.7478
5.2774 | -99.7413
5.2812 | -99.7478
5.2849 | -99.7661
5.2886 | -99.7814
5.2923 | -99.7901
5.2961 | -99.7963
5.2998 | -99.8063
5.3035 | -99.8134
5.3072 | -99.8266
5.311 | -99.8473
5.3147 | -99.8707
5.3184 | -99.8983
5.3222 | -99.9346
5.3259 | -99.9758
5.3296 | -100.0213
5.3333 | -100.0705
5.3371 | -100.1263
5.3408 | -100.1865
5.3445 | -100.2409
5.3482 | -100.2919
5.352 | -100.3315
5.3557 | -100.348
5.3594 | -100.3455
5.3631 | -100.3343
5.3669 | -100.3216
5.3706 | -100.3108
5.3743 | -100.3121
5.3781 | -100.3087
5.3818 | -100.2965
5.3855 | -100.2857
5.3892 | -100.2725
5.393 | -100.2548
5.3967 | -100.223
5.4004 | -100.1753
5.4041 | -100.1191
5.4079 | -100.0602
5.4116 | -99.9968
5.4153 | -99.9339
5.4191 | -99.8771
5.4228 | -99.8194
5.4265 | -99.7645
5.4302 | -99.7209
5.434 | -99.6914
5.4377 | -99.6814
5.4414 | -99.6924
5.4451 | -99.7204
5.4489 | -99.7637
5.4526 | -99.8198
5.4563 | -99.8855
5.4601 | -99.9513
5.4638 | -100.0019
5.4675 | -100.0308
5.4712 | -100.0308
5.475 | -100.0033
5.4787 | -99.9481
5.4824 | -99.882
5.4861 | -99.8135
5.4899 | -99.7483
5.4936 | -99.6904
5.4973 | -99.6533
5.501 | -99.6433
5.5048 | -99.656
5.5085 | -99.6909
5.5122 | -99.743
5.516 | -99.8013
5.5197 | -99.8643
5.5234 | -99.9161
5.5271 | -99.9542
5.5309 | -99.9802
5.5346 | -99.9879
5.5383 | -99.9825
5.542 | -99.964
5.5458 | -99.9446
5.5495 | -99.9335
5.5532 | -99.9302
5.557 | -99.934
5.5607 | -99.9443
5.5644 | -99.9613
5.5681 | -99.9785
5.5719 | -99.9976
5.5756 | -100.0133
5.5793 | -100.0346
5.583 | -100.0611
5.5868 | -100.0875
5.5905 | -100.115
5.5942 | -100.1428
5.5979 | -100.168
5.6017 | -100.1855
5.6054 | -100.1904
5.6091 | -100.1938
5.6129 | -100.1931
5.6166 | -100.1878
5.6203 | -100.1873
5.624 | -100.1842
5.6278 | -100.1839
5.6315 | -100.1805
5.6352 | -100.1813
5.6389 | -100.1797
5.6427 | -100.1807
5.6464 | -100.1932
5.6501 | -100.2151
5.6539 | -100.2363
5.6576 | -100.2578
5.6613 | -100.2771
5.665 | -100.284
5.6688 | -100.2787
5.6725 | -100.2691
5.6762 | -100.2525
5.6799 | -100.2201
5.6837 | -100.1759
5.6874 | -100.1327
5.6911 | -100.0938
5.6949 | -100.0567
5.6986 | -100.0231
5.7023 | -99.9969
5.706 | -99.9777
5.7098 | -99.9712
5.7135 | -99.975
5.7172 | -99.9896
5.7209 | -100.0087
5.7247 | -100.0326
5.7284 | -100.0573
5.7321 | -100.0787
5.7358 | -100.1039
5.7396 | -100.1326
5.7433 | -100.1517
5.747 | -100.1645
5.7508 | -100.1699
5.7545 | -100.1683
5.7582 | -100.1612
5.7619 | -100.1528
5.7657 | -100.149
5.7694 | -100.1428
5.7731 | -100.1307
5.7768 | -100.1188
5.7806 | -100.1134
5.7843 | -100.1189
5.788 | -100.1331
5.7918 | -100.1556
5.7955 | -100.1893
5.7992 | -100.2374
5.8029 | -100.3027
5.8067 | -100.382
5.8104 | -100.4683
5.8141 | -100.5557
5.8178 | -100.6281
5.8216 | -100.6737
5.8253 | -100.6996
5.829 | -100.6994
5.8328 | -100.6702
5.8365 | -100.6167
5.8402 | -100.5445
5.8439 | -100.4652
5.8477 | -100.3926
5.8514 | -100.3296
5.8551 | -100.28
5.8588 | -100.2413
5.8626 | -100.2177
5.8663 | -100.207
5.87 | -100.2108
5.8737 | -100.2244
5.8775 | -100.2479
5.8812 | -100.2764
5.8849 | -100.3021
5.8887 | -100.3248
5.8924 | -100.3492
5.8961 | -100.377
5.8998 | -100.4062
5.9036 | -100.4433
5.9073 | -100.4784
5.911 | -100.4993
5.9147 | -100.5211
5.9185 | -100.5414
5.9222 | -100.5545
5.9259 | -100.5613
5.9297 | -100.564
5.9334 | -100.5599
5.9371 | -100.5563
5.9408 | -100.5554
5.9446 | -100.5589
5.9483 | -100.5685
5.952 | -100.5757
5.9557 | -100.5875
5.9595 | -100.5966
5.9632 | -100.6087
5.9669 | -100.623
5.9707 | -100.6426
5.9744 | -100.6675
5.9781 | -100.6939
5.9818 | -100.7155
5.9856 | -100.7349
5.9893 | -100.7562
5.993 | -100.7724
5.9967 | -100.7825
6.0005 | -100.7791
6.0042 | -100.7673
6.0079 | -100.7471
6.0116 | -100.7183
6.0154 | -100.685
6.0191 | -100.6553
6.0228 | -100.634
6.0266 | -100.6238
6.0303 | -100.6213
6.034 | -100.6318
6.0377 | -100.6534
6.0415 | -100.6851
6.0452 | -100.7279
6.0489 | -100.7766
6.0526 | -100.8296
6.0564 | -100.8818
6.0601 | -100.9312
6.0638 | -100.9788
6.0676 | -101.0204
6.0713 | -101.0544
6.075 | -101.0869
6.0787 | -101.1115
6.0825 | -101.1292
6.0862 | -101.1408
6.0899 | -101.1406
6.0936 | -101.1276
6.0974 | -101.1051
6.1011 | -101.0769
6.1048 | -101.0378
6.1085 | -100.9975
6.1123 | -100.9504
6.116 | -100.8996
6.1197 | -100.8491
6.1235 | -100.8097
6.1272 | -100.7871
6.1309 | -100.7769
6.1346 | -100.7777
6.1384 | -100.7896
6.1421 | -100.813
6.1458 | -100.8438
6.1495 | -100.8785
6.1533 | -100.9118
6.157 | -100.941
6.1607 | -100.9684
6.1645 | -100.993
6.1682 | -101.0138
6.1719 | -101.0306
6.1756 | -101.0372
6.1794 | -101.0456
6.1831 | -101.0605
6.1868 | -101.0732
6.1905 | -101.0725
6.1943 | -101.0657
6.198 | -101.0614
6.2017 | -101.0519
6.2055 | -101.0396
6.2092 | -101.0271
6.2129 | -101.0116
6.2166 | -100.9952
6.2204 | -100.979
6.2241 | -100.9658
6.2278 | -100.9601
6.2315 | -100.9625
6.2353 | -100.9692
6.239 | -100.9824
6.2427 | -100.9948
6.2464 | -101.0067
6.2502 | -101.0146
6.2539 | -101.0137
6.2576 | -101.0024
6.2614 | -100.987
6.2651 | -100.9658
6.2688 | -100.939
6.2725 | -100.8973
6.2763 | -100.8456
6.28 | -100.7901
6.2837 | -100.7333
6.2874 | -100.6726
6.2912 | -100.6132
6.2949 | -100.5586
6.2986 | -100.5131
6.3024 | -100.4817
6.3061 | -100.4639
6.3098 | -100.4581
6.3135 | -100.467
6.3173 | -100.4989
6.321 | -100.5432
6.3247 | -100.5954
6.3284 | -100.6544
6.3322 | -100.7162
6.3359 | -100.7755
6.3396 | -100.8252
6.3434 | -100.8591
6.3471 | -100.8869
6.3508 | -100.9057
6.3545 | -100.9055
6.3583 | -100.9029
6.362 | -100.8996
6.3657 | -100.8916
6.3694 | -100.881
6.3732 | -100.8694
6.3769 | -100.8584
6.3806 | -100.8518
6.3843 | -100.8475
6.3881 | -100.8492
6.3918 | -100.8587
6.3955 | -100.8709
6.3993 | -100.8915
6.403 | -100.9192
6.4067 | -100.9531
6.4104 | -100.9795
6.4142 | -101.0109
6.4179 | -101.0523
6.4216 | -101.0962
6.4253 | -101.1398
6.4291 | -101.1902
6.4328 | -101.2407
6.4365 | -101.2937
6.4403 | -101.3411
6.444 | -101.3815
6.4477 | -101.4189
6.4514 | -101.4265
6.4552 | -101.4443
6.4589 | -101.4569
6.4626 | -101.4647
6.4663 | -101.4635
6.4701 | -101.452
6.4738 | -101.435
6.4775 | -101.4189
6.4812 | -101.4075
6.485 | -101.405
6.4887 | -101.4025
6.4924 | -101.4036
6.4962 | -101.4116
6.4999 | -101.4206
6.5036 | -101.4282
6.5073 | -101.4416
6.5111 | -101.4526
6.5148 | -101.4569
6.5185 | -101.4577
6.5222 | -101.4615
6.526 | -101.4631
6.5297 | -101.4627
6.5334 | -101.452
6.5372 | -101.4392
6.5409 | -101.4252
6.5446 | -101.4113
6.5483 | -101.393
6.5521 | -101.3797
6.5558 | -101.3708
6.5595 | -101.3562
6.5632 | -101.3534
6.567 | -101.3498
6.5707 | -101.3503
6.5744 | -101.3624
6.5781 | -101.3787
6.5819 | -101.3968
6.5856 | -101.4162
6.5893 | -101.4413
6.5931 | -101.4688
6.5968 | -101.4911
6.6005 | -101.5086
6.6042 | -101.5278
6.608 | -101.5445
6.6117 | -101.5521
6.6154 | -101.5566
6.6191 | -101.5639
6.6229 | -101.5655
6.6266 | -101.5648
6.6303 | -101.5686
6.6341 | -101.5766
6.6378 | -101.5812
6.6415 | -101.5843
6.6452 | -101.5895
6.649 | -101.5989
6.6527 | -101.6113
6.6564 | -101.6237
6.6601 | -101.6443
6.6639 | -101.6554
6.6676 | -101.663
6.6713 | -101.6591
6.6751 | -101.6345
6.6788 | -101.5955
6.6825 | -101.5387
6.6862 | -101.4673
6.69 | -101.3823
6.6937 | -101.2868
6.6974 | -101.1876
6.7011 | -101.0886
6.7049 | -100.9938
6.7086 | -100.9092
6.7123 | -100.841
6.716 | -100.7879
6.7198 | -100.747
6.7235 | -100.7259
6.7272 | -100.7233
6.731 | -100.7422
6.7347 | -100.7824
6.7384 | -100.8373
6.7421 | -100.9052
6.7459 | -100.9816
6.7496 | -101.0647
6.7533 | -101.1526
6.757 | -101.2387
6.7608 | -101.3194
6.7645 | -101.3938
6.7682 | -101.4571
6.772 | -101.5047
6.7757 | -101.538
6.7794 | -101.5595
6.7831 | -101.5784
6.7869 | -101.5902
6.7906 | -101.597
6.7943 | -101.5988
6.798 | -101.5965
6.8018 | -101.5935
6.8055 | -101.5924
6.8092 | -101.596
6.813 | -101.605
6.8167 | -101.6234
6.8204 | -101.6502
6.8241 | -101.6823
6.8279 | -101.7196
6.8316 | -101.7638
6.8353 | -101.8114
6.839 | -101.8585
6.8428 | -101.9056
6.8465 | -101.9483
6.8502 | -101.9801
6.8539 | -101.9949
6.8577 | -102.0026
6.8614 | -102.0029
6.8651 | -101.9922
6.8689 | -101.9709
6.8726 | -101.9419
6.8763 | -101.9071
6.88 | -101.8698
6.8838 | -101.8335
6.8875 | -101.7996
6.8912 | -101.7707
6.8949 | -101.7454
6.8987 | -101.7244
6.9024 | -101.7068
6.9061 | -101.693
6.9099 | -101.6774
6.9136 | -101.6604
6.9173 | -101.6476
6.921 | -101.6401
6.9248 | -101.6369
6.9285 | -101.6337
6.9322 | -101.6273
6.9359 | -101.6222
6.9397 | -101.6144
6.9434 | -101.5967
6.9471 | -101.5766
6.9508 | -101.5598
6.9546 | -101.5444
6.9583 | -101.5243
6.962 | -101.506
6.9658 | -101.4937
6.9695 | -101.49
6.9732 | -101.4903
6.9769 | -101.4972
6.9807 | -101.5083
6.9844 | -101.5287
6.9881 | -101.5881
6.9918 | -101.6239
6.9956 | -101.6689
6.9993 | -101.7211
7.003 | -101.7741
7.0068 | -101.8274
7.0105 | -101.8773
7.0142 | -101.9275
7.0179 | -101.9751
7.0217 | -102.0091
7.0254 | -102.0321
7.0291 | -102.0429
7.0328 | -102.0449
7.0366 | -102.0437
7.0403 | -102.0415
7.044 | -102.0425
7.0478 | -102.0442
7.0515 | -102.0441
7.0552 | -102.0419
7.0589 | -102.044
7.0627 | -102.0537
7.0664 | -102.061
7.0701 | -102.0668
7.0738 | -102.0738
7.0776 | -102.0804
7.0813 | -102.08
7.085 | -102.0761
7.0887 | -102.0671
7.0925 | -102.0531
7.0962 | -102.0382
7.0999 | -102.027
7.1037 | -102.0127
7.1074 | -102.0007
7.1111 | -101.9934
7.1148 | -101.9904
7.1186 | -101.9928
7.1223 | -101.9958
7.126 | -102.0017
7.1297 | -102.0125
7.1335 | -102.0262
7.1372 | -102.049
7.1409 | -102.0788
7.1447 | -102.1048
7.1484 | -102.1314
7.1521 | -102.1617
7.1558 | -102.1937
7.1596 | -102.2209
7.1633 | -102.245
7.167 | -102.2711
7.1707 | -102.2934
7.1745 | -102.3168
7.1782 | -102.3396
7.1819 | -102.3491
7.1857 | -102.3467
7.1894 | -102.3384
7.1931 | -102.3251
7.1968 | -102.2994
7.2006 | -102.2666
7.2043 | -102.2248
7.208 | -102.1725
7.2117 | -102.1124
7.2155 | -102.0498
7.2192 | -101.9886
7.2229 | -101.9307
7.2266 | -101.8803
7.2304 | -101.8383
7.2341 | -101.8033
7.2378 | -101.7789
7.2416 | -101.7687
7.2453 | -101.7698
7.249 | -101.7816
7.2527 | -101.809
7.2565 | -101.8507
7.2602 | -101.9055
7.2639 | -101.9728
7.2676 | -102.0432
7.2714 | -102.1165
7.2751 | -102.1884
7.2788 | -102.2588
7.2826 | -102.3222
7.2863 | -102.3763
7.29 | -102.4275
7.2937 | -102.4715
7.2975 | -102.5021
7.3012 | -102.51
7.3049 | -102.5106
7.3086 | -102.5017
7.3124 | -102.4863
7.3161 | -102.4682
7.3198 | -102.4457
7.3236 | -102.4195
7.3273 | -102.3955
7.331 | -102.3751
7.3347 | -102.3577
7.3385 | -102.3498
7.3422 | -102.3478
7.3459 | -102.3487
7.3496 | -102.3513
7.3534 | -102.3536
7.3571 | -102.3547
7.3608 | -102.3508
7.3645 | -102.3426
7.3683 | -102.3316
7.372 | -102.322
7.3757 | -102.3139
7.3795 | -102.3049
7.3832 | -102.2925
7.3869 | -102.2794
7.3906 | -102.2704
7.3944 | -102.2652
7.3981 | -102.2596
7.4018 | -102.2573
7.4055 | -102.2514
7.4093 | -102.2498
7.413 | -102.2433
7.4167 | -102.2285
7.4205 | -102.2139
7.4242 | -102.1959
7.4279 | -102.174
7.4316 | -102.1465
7.4354 | -102.1122
7.4391 | -102.0751
7.4428 | -102.0362
7.4465 | -101.9998
7.4503 | -101.966
7.454 | -101.9358
7.4577 | -101.9151
7.4614 | -101.9047
7.4652 | -101.9096
7.4689 | -101.9289
7.4726 | -101.9573
7.4764 | -101.9966
7.4801 | -102.045
7.4838 | -102.1039
7.4875 | -102.1712
7.4913 | -102.2472
7.495 | -102.3321
7.4987 | -102.4221
7.5024 | -102.5111
7.5062 | -102.5965
7.5099 | -102.6768
7.5136 | -102.7573
7.5174 | -102.829
7.5211 | -102.8903
7.5248 | -102.8942
7.5285 | -102.9238
7.5323 | -102.9333
7.536 | -102.9247
7.5397 | -102.9022
7.5434 | -102.8722
7.5472 | -102.8325
7.5509 | -102.7813
7.5546 | -102.7265
7.5584 | -102.6647
7.5621 | -102.6035
7.5658 | -102.544
7.5695 | -102.4865
7.5733 | -102.4385
7.577 | -102.3981
7.5807 | -102.3626
7.5844 | -102.3349
7.5882 | -102.3181
7.5919 | -102.3113
7.5956 | -102.311
7.5993 | -102.3165
7.6031 | -102.3217
7.6068 | -102.3256
7.6105 | -102.3294
7.6143 | -102.334
7.618 | -102.3348
7.6217 | -102.3261
7.6254 | -102.3094
7.6292 | -102.2877
7.6329 | -102.2621
7.6366 | -102.2378
7.6403 | -102.2133
7.6441 | -102.1874
7.6478 | -102.1634
7.6515 | -102.1395
7.6553 | -102.1181
7.659 | -102.1008
7.6627 | -102.0967
7.6664 | -102.1008
7.6702 | -102.1057
7.6739 | -102.1196
7.6776 | -102.138
7.6813 | -102.1622
7.6851 | -102.1914
7.6888 | -102.2203
7.6925 | -102.252
7.6963 | -102.2877
7.7 | -102.3276
7.7037 | -102.3632
7.7074 | -102.3953
7.7112 | -102.4238
7.7149 | -102.4454
7.7186 | -102.4629
7.7223 | -102.4751
7.7261 | -102.4816
7.7298 | -102.4874
7.7335 | -102.4904
7.7372 | -102.488
7.741 | -102.4769
7.7447 | -102.4547
7.7484 | -102.4236
7.7522 | -102.3862
7.7559 | -102.3451
7.7596 | -102.3015
7.7633 | -102.2574
7.7671 | -102.2112
7.7708 | -102.168
7.7745 | -102.1261
7.7782 | -102.0831
7.782 | -102.0489
7.7857 | -102.0228
7.7894 | -101.9979
7.7932 | -101.9781
7.7969 | -101.9638
7.8006 | -101.9474
7.8043 | -101.9377
7.8081 | -101.932
7.8118 | -101.9294
7.8155 | -101.9277
7.8192 | -101.9335
7.823 | -101.9462
7.8267 | -101.9627
7.8304 | -101.9785
7.8341 | -101.9981
7.8379 | -102.022
7.8416 | -102.043
7.8453 | -102.0597
7.8491 | -102.0792
7.8528 | -102.1001
7.8565 | -102.1277
7.8602 | -102.1603
7.864 | -102.1921
7.8677 | -102.2244
7.8714 | -102.2552
7.8751 | -102.2826
7.8789 | -102.3106
7.8826 | -102.3424
7.8863 | -102.3681
7.8901 | -102.3921
7.8938 | -102.4161
7.8975 | -102.4349
7.9012 | -102.4537
7.905 | -102.4722
7.9087 | -102.486
7.9124 | -102.4987
7.9161 | -102.514
7.9199 | -102.5188
7.9236 | -102.5201
7.9273 | -102.5193
7.9311 | -102.5145
7.9348 | -102.5017
7.9385 | -102.4877
7.9422 | -102.4793
7.946 | -102.4733
7.9497 | -102.4637
7.9534 | -102.4559
7.9571 | -102.4566
7.9609 | -102.4653
7.9646 | -102.4774
7.9683 | -102.4951
7.972 | -102.5158
7.9758 | -102.544
7.9795 | -102.5757
7.9832 | -102.6033
7.987 | -102.6321
7.9907 | -102.667
7.9944 | -102.6977
7.9981 | -102.7231
8.0019 | -102.7463
8.0056 | -102.7622
8.0093 | -102.7751
8.013 | -102.7859
8.0168 | -102.7909
8.0205 | -102.7918
8.0242 | -102.7967
8.028 | -102.8
8.0317 | -102.8034
8.0354 | -102.8143
8.0391 | -102.828
8.0429 | -102.8463
8.0466 | -102.8703
8.0503 | -102.8927
8.054 | -102.9169
8.0578 | -102.9414
8.0615 | -102.969
8.0652 | -102.9826
8.0689 | -103.0003
8.0727 | -103.0102
8.0764 | -103.0167
8.0801 | -103.019
8.0839 | -103.0169
8.0876 | -103.0078
8.0913 | -102.9912
8.095 | -102.9704
8.0988 | -102.9504
8.1025 | -102.926
8.1062 | -102.8993
8.1099 | -102.8675
8.1137 | -102.8376
8.1174 | -102.8111
8.1211 | -102.7892
8.1249 | -102.772
8.1286 | -102.7539
8.1323 | -102.7391
8.136 | -102.7317
8.1398 | -102.734
8.1435 | -102.7384
8.1472 | -102.7469
8.1509 | -102.7614
8.1547 | -102.775
8.1584 | -102.786
8.1621 | -102.8002
8.1659 | -102.8154
8.1696 | -102.835
8.1733 | -102.8491
8.177 | -102.8627
8.1808 | -102.8752
8.1845 | -102.8864
8.1882 | -102.892
8.1919 | -102.8916
8.1957 | -102.8885
8.1994 | -102.8797
8.2031 | -102.8648
8.2068 | -102.8499
8.2106 | -102.8372
8.2143 | -102.8214
8.218 | -102.8011
8.2218 | -102.78
8.2255 | -102.7624
8.2292 | -102.7419
8.2329 | -102.7192
8.2367 | -102.696
8.2404 | -102.6731
8.2441 | -102.6508
8.2478 | -102.6316
8.2516 | -102.6194
8.2553 | -102.6053
8.259 | -102.5882
8.2628 | -102.5703
8.2665 | -102.5573
8.2702 | -102.5485
8.2739 | -102.5442
8.2777 | -102.5442
8.2814 | -102.5486
8.2851 | -102.5519
8.2888 | -102.5574
8.2926 | -102.5625
8.2963 | -102.5701
8.3 | -102.5813
8.3038 | -102.5949
8.3075 | -102.6097
8.3112 | -102.6237
8.3149 | -102.6383
8.3187 | -102.6545
8.3224 | -102.6728
8.3261 | -102.6919
8.3298 | -102.7086
8.3336 | -102.7264
8.3373 | -102.7536
8.341 | -102.7869
8.3447 | -102.8202
8.3485 | -102.8534
8.3522 | -102.8886
8.3559 | -102.9262
8.3597 | -102.9588
8.3634 | -102.988
8.3671 | -103.0134
8.3708 | -103.0355
8.3746 | -103.0538
8.3783 | -103.0712
8.382 | -103.085
8.3857 | -103.087
8.3895 | -103.0876
8.3932 | -103.0875
8.3969 | -103.0812
8.4007 | -103.0719
8.4044 | -103.0626
8.4081 | -103.052
8.4118 | -103.0379
8.4156 | -103.027
8.4193 | -103.017
8.423 | -103.0034
8.4267 | -102.9963
8.4305 | -102.9904
8.4342 | -102.9912
8.4379 | -102.9968
8.4416 | -102.9964
8.4454 | -102.9954
8.4491 | -102.9989
8.4528 | -103.0028
8.4566 | -103.0027
8.4603 | -102.9934
8.464 | -102.9844
8.4677 | -102.9772
8.4715 | -102.9619
8.4752 | -102.9447
8.4789 | -102.9224
8.4826 | -102.8952
8.4864 | -102.8651
8.4901 | -102.8343
8.4938 | -102.7984
8.4976 | -102.7593
8.5013 | -102.7217
8.505 | -102.6842
8.5087 | -102.6517
8.5125 | -102.6192
8.5162 | -102.5926
8.5199 | -102.5739
8.5236 | -102.5619
8.5274 | -102.5556
8.5311 | -102.555
8.5348 | -102.5641
8.5386 | -102.5786
8.5423 | -102.5983
8.546 | -102.6253
8.5497 | -102.6551
8.5535 | -102.6877
8.5572 | -102.7177
8.5609 | -102.7423
8.5646 | -102.7707
8.5684 | -102.7974
8.5721 | -102.8225
8.5758 | -102.8472
8.5795 | -102.8688
8.5833 | -102.8893
8.587 | -102.9067
8.5907 | -102.9211
8.5945 | -102.938
8.5982 | -102.9468
8.6019 | -102.9462
8.6056 | -102.9495
8.6094 | -102.9478
8.6131 | -102.9399
8.6168 | -102.9273
8.6205 | -102.9162
8.6243 | -102.9032
8.628 | -102.8873
8.6317 | -102.8735
8.6355 | -102.8607
8.6392 | -102.8459
8.6429 | -102.8275
8.6466 | -102.811
8.6504 | -102.7972
8.6541 | -102.7857
8.6578 | -102.7763
8.6615 | -102.7676
8.6653 | -102.7641
8.669 | -102.7653
8.6727 | -102.7717
8.6765 | -102.7806
8.6802 | -102.793
8.6839 | -102.8099
8.6876 | -102.8278
8.6914 | -102.8407
8.6951 | -102.852
8.6988 | -102.8574
8.7025 | -102.8573
8.7063 | -102.8557
8.71 | -102.85
8.7137 | -102.8407
8.7174 | -102.8325
8.7212 | -102.8294
8.7249 | -102.8287
8.7286 | -102.8367
8.7324 | -102.8498
8.7361 | -102.8707
8.7398 | -102.902
8.7435 | -102.9396
8.7473 | -102.9804
8.751 | -103.0273
8.7547 | -103.0799
8.7584 | -103.1355
8.7622 | -103.1906
8.7659 | -103.2462
8.7696 | -103.3027
8.7734 | -103.3595
8.7771 | -103.4135
8.7808 | -103.4651
8.7845 | -103.5121
8.7883 | -103.5535
8.792 | -103.5888
8.7957 | -103.6157
8.7994 | -103.6406
8.8032 | -103.656
8.8069 | -103.6625
8.8106 | -103.6661
8.8143 | -103.6659
8.8181 | -103.663
8.8218 | -103.6592
8.8255 | -103.649
8.8293 | -103.634
8.833 | -103.62
8.8367 | -103.6046
8.8404 | -103.5872
8.8442 | -103.5677
8.8479 | -103.5428
8.8516 | -103.5133
8.8553 | -103.4871
8.8591 | -103.4558
8.8628 | -103.4238
8.8665 | -103.4
8.8703 | -103.378
8.874 | -103.3589
8.8777 | -103.3432
8.8814 | -103.3266
8.8852 | -103.315
8.8889 | -103.3076
8.8926 | -103.3036
8.8963 | -103.302
8.9001 | -103.3015
8.9038 | -103.3026
8.9075 | -103.3054
8.9113 | -103.308
8.915 | -103.3127
8.9187 | -103.3193
8.9224 | -103.3268
8.9262 | -103.3343
8.9299 | -103.339
8.9336 | -103.3508
8.9373 | -103.3696
8.9411 | -103.3863
8.9448 | -103.399
8.9485 | -103.4124
8.9522 | -103.4319
8.956 | -103.4575
8.9597 | -103.4867
8.9634 | -103.5118
8.9672 | -103.5343
8.9709 | -103.5599
8.9746 | -103.5856
8.9783 | -103.6134
8.9821 | -103.6389
8.9858 | -103.6578
8.9895 | -103.675
8.9932 | -103.689
8.997 | -103.6988
9.0007 | -103.7009
9.0044 | -103.6984
9.0082 | -103.6952
9.0119 | -103.6921
9.0156 | -103.6846
9.0193 | -103.6693
9.0231 | -103.6493
9.0268 | -103.6275
9.0305 | -103.6002
9.0342 | -103.5692
9.038 | -103.5381
9.0417 | -103.5124
9.0454 | -103.4857
9.0491 | -103.456
9.0529 | -103.4324
9.0566 | -103.4148
9.0603 | -103.4007
9.0641 | -103.3891
9.0678 | -103.3838
9.0715 | -103.3847
9.0752 | -103.3882
9.079 | -103.3918
9.0827 | -103.3955
9.0864 | -103.3992
9.0901 | -103.4039
9.0939 | -103.4105
9.0976 | -103.4168
9.1013 | -103.4219
9.1051 | -103.4278
9.1088 | -103.4342
9.1125 | -103.4413
9.1162 | -103.4492
9.12 | -103.4553
9.1237 | -103.4615
9.1274 | -103.4662
9.1311 | -103.4778
9.1349 | -103.4856
9.1386 | -103.4879
9.1423 | -103.4909
9.1461 | -103.4921
9.1498 | -103.4914
9.1535 | -103.4955
9.1572 | -103.4987
9.161 | -103.4998
9.1647 | -103.5061
9.1684 | -103.5107
9.1721 | -103.5116
9.1759 | -103.5112
9.1796 | -103.5167
9.1833 | -103.5304
9.187 | -103.5397
9.1908 | -103.5473
9.1945 | -103.5574
9.1982 | -103.5652
9.202 | -103.5664
9.2057 | -103.5704
9.2094 | -103.5735
9.2131 | -103.5759
9.2169 | -103.574
9.2206 | -103.5711
9.2243 | -103.5679
9.228 | -103.5662
9.2318 | -103.5641
9.2355 | -103.558
9.2392 | -103.5508
9.243 | -103.5443
9.2467 | -103.5437
9.2504 | -103.5456
9.2541 | -103.5479
9.2579 | -103.5507
9.2616 | -103.5555
9.2653 | -103.5588
9.269 | -103.5663
9.2728 | -103.5756
9.2765 | -103.586
9.2802 | -103.5978
9.284 | -103.6053
9.2877 | -103.6126
9.2914 | -103.6157
9.2951 | -103.6133
9.2989 | -103.6107
9.3026 | -103.6057
9.3063 | -103.5991
9.31 | -103.5906
9.3138 | -103.5753
9.3175 | -103.5521
9.3212 | -103.5235
9.3249 | -103.4916
9.3287 | -103.456
9.3324 | -103.4171
9.3361 | -103.3774
9.3399 | -103.3425
9.3436 | -103.3047
9.3473 | -103.2656
9.351 | -103.2282
9.3548 | -103.1989
9.3585 | -103.1733
9.3622 | -103.1507
9.3659 | -103.1325
9.3697 | -103.1202
9.3734 | -103.1143
9.3771 | -103.1142
9.3809 | -103.1243
9.3846 | -103.1386
9.3883 | -103.1597
9.392 | -103.1859
9.3958 | -103.2156
9.3995 | -103.2501
9.4032 | -103.2931
9.4069 | -103.3356
9.4107 | -103.3735
9.4144 | -103.4104
9.4181 | -103.4474
9.4218 | -103.4835
9.4256 | -103.5151
9.4293 | -103.5413
9.433 | -103.5658
9.4368 | -103.5791
9.4405 | -103.5839
9.4442 | -103.5819
9.4479 | -103.5722
9.4517 | -103.5594
9.4554 | -103.5452
9.4591 | -103.5302
9.4628 | -103.5138
9.4666 | -103.4953
9.4703 | -103.4768
9.474 | -103.4604
9.4778 | -103.4443
9.4815 | -103.4266
9.4852 | -103.4165
9.4889 | -103.4112
9.4927 | -103.4081
9.4964 | -103.4095
9.5001 | -103.4159
9.5038 | -103.4265
9.5076 | -103.4399
9.5113 | -103.4557
9.515 | -103.4733
9.5188 | -103.4846
9.5225 | -103.4978
9.5262 | -103.5164
9.5299 | -103.5399
9.5337 | -103.5661
9.5374 | -103.5927
9.5411 | -103.6188
9.5448 | -103.6489
9.5486 | -103.6846
9.5523 | -103.7196
9.556 | -103.7523
9.5597 | -103.7885
9.5635 | -103.8251
9.5672 | -103.862
9.5709 | -103.8952
9.5747 | -103.9279
9.5784 | -103.9544
9.5821 | -103.9806
9.5858 | -104.0063
9.5896 | -104.025
9.5933 | -104.038
9.597 | -104.0469
9.6007 | -104.0529
9.6045 | -104.0555
9.6082 | -104.0549
9.6119 | -104.0523
9.6157 | -104.0439
9.6194 | -104.0362
9.6231 | -104.0233
9.6268 | -104.0063
9.6306 | -103.9863
9.6343 | -103.9669
9.638 | -103.9491
9.6417 | -103.926
9.6455 | -103.9033
9.6492 | -103.8841
9.6529 | -103.8626
9.6567 | -103.8458
9.6604 | -103.8315
9.6641 | -103.8137
9.6678 | -103.7994
9.6716 | -103.7882
9.6753 | -103.7882
9.679 | -103.78
9.6827 | -103.7755
9.6865 | -103.7765
9.6902 | -103.7806
9.6939 | -103.7852
9.6976 | -103.7873
9.7014 | -103.7876
9.7051 | -103.7894
9.7088 | -103.7941
9.7126 | -103.8011
9.7163 | -103.8043
9.72 | -103.8094
9.7237 | -103.8073
9.7275 | -103.8031
9.7312 | -103.7996
9.7349 | -103.7952
9.7386 | -103.7841
9.7424 | -103.7716
9.7461 | -103.7609
9.7498 | -103.7455
9.7536 | -103.7256
9.7573 | -103.7063
9.761 | -103.6885
9.7647 | -103.6713
9.7685 | -103.6544
9.7722 | -103.6391
9.7759 | -103.6287
9.7796 | -103.6157
9.7834 | -103.6089
9.7871 | -103.6081
9.7908 | -103.6115
9.7945 | -103.6194
9.7983 | -103.624
9.802 | -103.6326
9.8057 | -103.6499
9.8095 | -103.6677
9.8132 | -103.6844
9.8169 | -103.7037
9.8206 | -103.7269
9.8244 | -103.7515
9.8281 | -103.7745
9.8318 | -103.7982
9.8355 | -103.8228
9.8393 | -103.8475
9.843 | -103.8727
9.8467 | -103.897
9.8505 | -103.9213
9.8542 | -103.9452
9.8579 | -103.9695
9.8616 | -103.9969
9.8654 | -104.0231
9.8691 | -104.0508
9.8728 | -104.0755
9.8765 | -104.1002
9.8803 | -104.1205
9.884 | -104.1398
9.8877 | -104.1524
9.8915 | -104.1585
9.8952 | -104.1671
9.8989 | -104.174
9.9026 | -104.1693
9.9064 | -104.1568
9.9101 | -104.1415
9.9138 | -104.1188
9.9175 | -104.092
9.9213 | -104.0585
9.925 | -104.0226
9.9287 | -103.9826
9.9324 | -103.9372
9.9362 | -103.8815
9.9399 | -103.8272
9.9436 | -103.78
9.9474 | -103.7312
9.9511 | -103.679
9.9548 | -103.6274
9.9585 | -103.578
9.9623 | -103.5279
9.966 | -103.4791
9.9697 | -103.4344
9.9734 | -103.3936
9.9772 | -103.3556
9.9809 | -103.3213
9.9846 | -103.2933
9.9884 | -103.2697
9.9921 | -103.2503
9.9958 | -103.233
9.9995 | -103.2177
10.0033 | -103.2108
10.007 | -103.2096
10.0107 | -103.2102
10.0144 | -103.2159
10.0182 | -103.2277
10.0219 | -103.2448
10.0256 | -103.2683
10.0294 | -103.2988
10.0331 | -103.3346
10.0368 | -103.3712
10.0405 | -103.4081
10.0443 | -103.4518
10.048 | -103.5005
10.0517 | -103.5521
10.0554 | -103.604
10.0592 | -103.6586
10.0629 | -103.7121
10.0666 | -103.7648
10.0703 | -103.8181
10.0741 | -103.8645
10.0778 | -103.9077
10.0815 | -103.9567
10.0853 | -104.0003
10.089 | -104.0384
10.0927 | -104.0723
10.0964 | -104.1025
10.1002 | -104.1292
10.1039 | -104.1546
10.1076 | -104.1787
10.1113 | -104.1978
10.1151 | -104.2081
10.1188 | -104.2093
10.1225 | -104.2107
10.1263 | -104.2126
10.13 | -104.2097
10.1337 | -104.2012
10.1374 | -104.1909
10.1412 | -104.1787
10.1449 | -104.1677
10.1486 | -104.152
10.1523 | -104.1331
10.1561 | -104.1163
10.1598 | -104.0997
10.1635 | -104.0791
10.1672 | -104.0595
10.171 | -104.04
10.1747 | -104.0229
10.1784 | -104.0078
10.1822 | -103.9937
10.1859 | -103.9787
10.1896 | -103.9626
10.1933 | -103.9507
10.1971 | -103.9374
10.2008 | -103.9259
10.2045 | -103.919
10.2082 | -103.9133
10.212 | -103.9092
10.2157 | -103.9092
10.2194 | -103.906
10.2232 | -103.9044
10.2269 | -103.9003
10.2306 | -103.8964
10.2343 | -103.8923
10.2381 | -103.8884
10.2418 | -103.8861
10.2455 | -103.8877
10.2492 | -103.8892
10.253 | -103.8894
10.2567 | -103.8865
10.2604 | -103.8833
10.2642 | -103.8874
10.2679 | -103.8965
10.2716 | -103.9024
10.2753 | -103.909
10.2791 | -103.9169
10.2828 | -103.928
10.2865 | -103.94
10.2902 | -103.9513
10.294 | -103.9639
10.2977 | -103.978
10.3014 | -103.9945
10.3051 | -104.0089
10.3089 | -104.0225
10.3126 | -104.0394
10.3163 | -104.053
10.3201 | -104.0679
10.3238 | -104.0821
10.3275 | -104.0984
10.3312 | -104.1183
10.335 | -104.138
10.3387 | -104.1547
10.3424 | -104.17
10.3461 | -104.1836
10.3499 | -104.1941
10.3536 | -104.2084
10.3573 | -104.2194
10.3611 | -104.2284
10.3648 | -104.2364
10.3685 | -104.2411
10.3722 | -104.2479
10.376 | -104.2583
10.3797 | -104.2686
10.3834 | -104.2726
10.3871 | -104.2736
10.3909 | -104.2761
10.3946 | -104.2777
10.3983 | -104.2762
10.402 | -104.2717
10.4058 | -104.257
10.4095 | -104.2441
10.4132 | -104.2315
10.417 | -104.2125
10.4207 | -104.1907
10.4244 | -104.1685
10.4281 | -104.1444
10.4319 | -104.1198
10.4356 | -104.0955
10.4393 | -104.0675
10.443 | -104.0373
10.4468 | -104.0052
10.4505 | -103.9749
10.4542 | -103.9478
10.458 | -103.9216
10.4617 | -103.8976
10.4654 | -103.8753
10.4691 | -103.8577
10.4729 | -103.8398
10.4766 | -103.8219
10.4803 | -103.8085
10.484 | -103.8012
10.4878 | -103.8002
10.4915 | -103.8046
10.4952 | -103.812
10.499 | -103.8213
10.5027 | -103.8335
10.5064 | -103.8426
10.5101 | -103.8557
10.5139 | -103.8721
10.5176 | -103.8877
10.5213 | -103.9004
10.525 | -103.9119
10.5288 | -103.9203
10.5325 | -103.926
10.5362 | -103.9348
10.5399 | -103.9466
10.5437 | -103.9532
10.5474 | -103.9612
10.5511 | -103.9702
10.5549 | -103.9798
10.5586 | -103.9934
10.5623 | -104.0011
10.566 | -104.0044
10.5698 | -104.0064
10.5735 | -104.0062
10.5772 | -104.0018
10.5809 | -103.9977
10.5847 | -103.9974
10.5884 | -103.9927
10.5921 | -103.9854
10.5959 | -103.979
10.5996 | -103.975
10.6033 | -103.97
10.607 | -103.9641
10.6108 | -103.9596
10.6145 | -103.9577
10.6182 | -103.9571
10.6219 | -103.9593
10.6257 | -103.9681
10.6294 | -103.9793
10.6331 | -103.9888
10.6369 | -103.9981
10.6406 | -104.0097
10.6443 | -104.0251
10.648 | -104.0418
10.6518 | -104.0599
10.6555 | -104.0769
10.6592 | -104.0937
10.6629 | -104.1119
10.6667 | -104.1279
10.6704 | -104.142
10.6741 | -104.1568
10.6778 | -104.1671
10.6816 | -104.175
10.6853 | -104.1805
10.689 | -104.1803
10.6928 | -104.1792
10.6965 | -104.1766
10.7002 | -104.1684
10.7039 | -104.1587
10.7077 | -104.1442
10.7114 | -104.1254
10.7151 | -104.1058
10.7188 | -104.0864
10.7226 | -104.0618
10.7263 | -104.0333
10.73 | -104.0035
10.7338 | -103.9725
10.7375 | -103.9389
10.7412 | -103.9102
10.7449 | -103.8859
10.7487 | -103.8627
10.7524 | -103.8521
10.7561 | -103.8315
10.7598 | -103.8148
10.7636 | -103.8042
10.7673 | -103.7995
10.771 | -103.7969
10.7747 | -103.7984
10.7785 | -103.8029
10.7822 | -103.8112
10.7859 | -103.8236
10.7897 | -103.8399
10.7934 | -103.857
10.7971 | -103.8743
10.8008 | -103.891
10.8046 | -103.9124
10.8083 | -103.9359
10.812 | -103.9569
10.8157 | -103.9777
10.8195 | -103.9991
10.8232 | -104.0217
10.8269 | -104.0461
10.8307 | -104.0724
10.8344 | -104.0971
10.8381 | -104.1174
10.8418 | -104.1358
10.8456 | -104.1565
10.8493 | -104.1733
10.853 | -104.1881
10.8567 | -104.2027
10.8605 | -104.2159
10.8642 | -104.2322
10.8679 | -104.2438
10.8717 | -104.2532
10.8754 | -104.2624
10.8791 | -104.2758
10.8828 | -104.2926
10.8866 | -104.3112
10.8903 | -104.3261
10.894 | -104.3383
10.8977 | -104.3516
10.9015 | -104.3656
10.9052 | -104.379
10.9089 | -104.3912
10.9126 | -104.4041
10.9164 | -104.4177
10.9201 | -104.4259
10.9238 | -104.4287
10.9276 | -104.4321
10.9313 | -104.4333
10.935 | -104.4306
10.9387 | -104.4253
10.9425 | -104.4169
10.9462 | -104.4061
10.9499 | -104.3949
10.9536 | -104.3811
10.9574 | -104.3646
10.9611 | -104.3468
10.9648 | -104.3227
10.9686 | -104.2986
10.9723 | -104.2773
10.976 | -104.2561
10.9797 | -104.2355
10.9835 | -104.2156
10.9872 | -104.1945
10.9909 | -104.1734
10.9946 | -104.1548
10.9984 | -104.143
11.0021 | -104.1325
11.0058 | -104.1225
11.0096 | -104.1159
11.0133 | -104.1154
11.017 | -104.1174
11.0207 | -104.123
11.0245 | -104.1303
11.0282 | -104.1437
11.0319 | -104.1572
11.0356 | -104.1665
11.0394 | -104.1802
11.0431 | -104.1975
11.0468 | -104.2148
11.0505 | -104.2351
11.0543 | -104.2589
11.058 | -104.2806
11.0617 | -104.3012
11.0655 | -104.3203
11.0692 | -104.3394
11.0729 | -104.3553
11.0766 | -104.3657
11.0804 | -104.3763
11.0841 | -104.3866
11.0878 | -104.3889
11.0915 | -104.3871
11.0953 | -104.386
11.099 | -104.381
11.1027 | -104.3742
11.1065 | -104.3649
11.1102 | -104.353
11.1139 | -104.3404
11.1176 | -104.328
11.1214 | -104.3177
11.1251 | -104.3108
11.1288 | -104.3013
11.1325 | -104.29
11.1363 | -104.2779
11.14 | -104.2656
11.1437 | -104.2555
11.1474 | -104.2506
11.1512 | -104.2472
11.1549 | -104.2418
11.1586 | -104.2345
11.1624 | -104.23
11.1661 | -104.2233
11.1698 | -104.2125
11.1735 | -104.2007
11.1773 | -104.1921
11.181 | -104.1807
11.1847 | -104.1694
11.1884 | -104.1577
11.1922 | -104.1464
11.1959 | -104.1401
11.1996 | -104.1315
11.2034 | -104.1208
11.2071 | -104.1115
11.2108 | -104.1017
11.2145 | -104.0946
11.2183 | -104.0892
11.222 | -104.0855
11.2257 | -104.0863
11.2294 | -104.086
11.2332 | -104.0885
11.2369 | -104.0949
11.2406 | -104.1028
11.2444 | -104.1137
11.2481 | -104.1284
11.2518 | -104.1474
11.2555 | -104.1664
11.2593 | -104.1912
11.263 | -104.216
11.2667 | -104.2437
11.2704 | -104.2746
11.2742 | -104.3055
11.2779 | -104.3357
11.2816 | -104.3637
11.2853 | -104.3919
11.2891 | -104.416
11.2928 | -104.4438
11.2965 | -104.4685
11.3003 | -104.4896
11.304 | -104.5097
11.3077 | -104.5259
11.3114 | -104.5332
11.3152 | -104.5362
11.3189 | -104.539
11.3226 | -104.5376
11.3263 | -104.5335
11.3301 | -104.5273
11.3338 | -104.5214
11.3375 | -104.512
11.3413 | -104.4975
11.345 | -104.4829
11.3487 | -104.4681
11.3524 | -104.4524
11.3562 | -104.4342
11.3599 | -104.4151
11.3636 | -104.3941
11.3673 | -104.3753
11.3711 | -104.3577
11.3748 | -104.3425
11.3785 | -104.3237
11.3823 | -104.3047
11.386 | -104.2891
11.3897 | -104.2729
11.3934 | -104.2603
11.3972 | -104.25
11.4009 | -104.2385
11.4046 | -104.2232
11.4083 | -104.2062
11.4121 | -104.1904
11.4158 | -104.1709
11.4195 | -104.1553
11.4232 | -104.1394
11.427 | -104.126
11.4307 | -104.1139
11.4344 | -104.0994
11.4382 | -104.0864
11.4419 | -104.071
11.4456 | -104.0555
11.4493 | -104.0388
11.4531 | -104.0233
11.4568 | -104.011
11.4605 | -103.9994
11.4642 | -103.9922
11.468 | -103.989
11.4717 | -103.9842
11.4754 | -103.9845
11.4792 | -103.9901
11.4829 | -103.9972
11.4866 | -104.004
11.4903 | -104.0152
11.4941 | -104.0284
11.4978 | -104.0432
11.5015 | -104.0618
11.5052 | -104.0794
11.509 | -104.1009
11.5127 | -104.1253
11.5164 | -104.1455
11.5201 | -104.1663
11.5239 | -104.1909
11.5276 | -104.216
11.5313 | -104.2418
11.5351 | -104.2647
11.5388 | -104.2838
11.5425 | -104.3047
11.5462 | -104.3264
11.55 | -104.3439
11.5537 | -104.3564
11.5574 | -104.3677
11.5611 | -104.3729
11.5649 | -104.3754
11.5686 | -104.3792
11.5723 | -104.3828
11.5761 | -104.3848
11.5798 | -104.3877
11.5835 | -104.3911
11.5872 | -104.3912
11.591 | -104.3895
11.5947 | -104.3865
11.5984 | -104.3841
11.6021 | -104.3869
11.6059 | -104.394
11.6096 | -104.4041
11.6133 | -104.4121
11.6171 | -104.4223
11.6208 | -104.4316
11.6245 | -104.4413
11.6282 | -104.4528
11.632 | -104.4629
11.6357 | -104.4724
11.6394 | -104.4809
11.6431 | -104.4857
11.6469 | -104.494
11.6506 | -104.501
11.6543 | -104.5046
11.658 | -104.5062
11.6618 | -104.501
11.6655 | -104.4929
11.6692 | -104.4827
11.673 | -104.4717
11.6767 | -104.4548
11.6804 | -104.4361
11.6841 | -104.4202
11.6879 | -104.4008
11.6916 | -104.3781
11.6953 | -104.3555
11.699 | -104.3336
11.7028 | -104.3162
11.7065 | -104.3007
11.7102 | -104.2838
11.714 | -104.2692
11.7177 | -104.259
11.7214 | -104.2533
11.7251 | -104.2507
11.7289 | -104.254
11.7326 | -104.2628
11.7363 | -104.2728
11.74 | -104.2873
11.7438 | -104.3072
11.7475 | -104.3257
11.7512 | -104.3507
11.7549 | -104.3773
11.7587 | -104.4064
11.7624 | -104.4374
11.7661 | -104.4693
11.7699 | -104.4991
11.7736 | -104.5277
11.7773 | -104.5572
11.781 | -104.5871
11.7848 | -104.6158
11.7885 | -104.643
11.7922 | -104.6674
11.7959 | -104.6907
11.7997 | -104.7119
11.8034 | -104.7302
11.8071 | -104.7455
11.8109 | -104.759
11.8146 | -104.7632
11.8183 | -104.7651
11.822 | -104.7647
11.8258 | -104.7449
11.8295 | -104.7353
11.8332 | -104.7237
11.8369 | -104.71
11.8407 | -104.695
11.8444 | -104.6749
11.8481 | -104.6541
11.8519 | -104.6303
11.8556 | -104.6025
11.8593 | -104.5728
11.863 | -104.5416
11.8668 | -104.506
11.8705 | -104.4732
11.8742 | -104.4407
11.8779 | -104.4101
11.8817 | -104.3817
11.8854 | -104.3577
11.8891 | -104.3337
11.8928 | -104.3064
11.8966 | -104.2827
11.9003 | -104.2635
11.904 | -104.243
11.9078 | -104.222
11.9115 | -104.2001
11.9152 | -104.1809
11.9189 | -104.1669
11.9227 | -104.1531
11.9264 | -104.1433
11.9301 | -104.1352
11.9338 | -104.1284
11.9376 | -104.1247
11.9413 | -104.1246
11.945 | -104.1231
11.9488 | -104.1223
11.9525 | -104.1235
11.9562 | -104.125
11.9599 | -104.1303
11.9637 | -104.1381
11.9674 | -104.1495
11.9711 | -104.1624
11.9748 | -104.1777
11.9786 | -104.1963
11.9823 | -104.2159
11.986 | -104.2379
11.9898 | -104.2626
11.9935 | -104.2906
11.9972 | -104.3205
12.0009 | -104.355
12.0047 | -104.3892
12.0084 | -104.4226
12.0121 | -104.4571
12.0158 | -104.4892
12.0196 | -104.5262
12.0233 | -104.5633
12.027 | -104.5988
12.0307 | -104.6357
12.0345 | -104.6709
12.0382 | -104.703
12.0419 | -104.7353
12.0457 | -104.7641
12.0494 | -104.7919
12.0531 | -104.8158
12.0568 | -104.8362
12.0606 | -104.853
12.0643 | -104.8701
12.068 | -104.8877
12.0717 | -104.9031
12.0755 | -104.9138
12.0792 | -104.9254
12.0829 | -104.9421
12.0867 | -104.9612
12.0904 | -104.976
12.0941 | -104.9883
12.0978 | -105.0008
12.1016 | -105.0145
12.1053 | -105.0292
12.109 | -105.0463
12.1127 | -105.0622
12.1165 | -105.0771
12.1202 | -105.0944
12.1239 | -105.1122
12.1276 | -105.1261
12.1314 | -105.1369
12.1351 | -105.1431
12.1388 | -105.1507
12.1426 | -105.156
12.1463 | -105.1625
12.15 | -105.1664
12.1537 | -105.1693
12.1575 | -105.1687
12.1612 | -105.166
12.1649 | -105.1589
12.1686 | -105.1479
12.1724 | -105.136
12.1761 | -105.1196
12.1798 | -105.1039
12.1836 | -105.0896
12.1873 | -105.0694
12.191 | -105.0494
12.1947 | -105.0267
12.1985 | -104.9989
12.2022 | -104.9728
12.2059 | -104.9458
12.2096 | -104.9206
12.2134 | -104.8951
12.2171 | -104.8705
12.2208 | -104.8461
12.2246 | -104.82
12.2283 | -104.7952
12.232 | -104.7704
12.2357 | -104.7475
12.2395 | -104.724
12.2432 | -104.6976
12.2469 | -104.6716
12.2506 | -104.6491
12.2544 | -104.6291
12.2581 | -104.6108
12.2618 | -104.5925
12.2655 | -104.5733
12.2693 | -104.5593
12.273 | -104.5473
12.2767 | -104.5348
12.2805 | -104.5229
12.2842 | -104.5113
12.2879 | -104.5031
12.2916 | -104.4935
12.2954 | -104.4867
12.2991 | -104.4804
12.3028 | -104.4755
12.3065 | -104.4744
12.3103 | -104.4763
12.314 | -104.4775
12.3177 | -104.4787
12.3215 | -104.4802
12.3252 | -104.4832
12.3289 | -104.4842
12.3326 | -104.4901
12.3364 | -104.4995
12.3401 | -104.511
12.3438 | -104.5227
12.3475 | -104.5358
12.3513 | -104.5511
12.355 | -104.5656
12.3587 | -104.5789
12.3625 | -104.594
12.3662 | -104.6067
12.3699 | -104.6212
12.3736 | -104.635
12.3774 | -104.6503
12.3811 | -104.6647
12.3848 | -104.679
12.3885 | -104.69
12.3923 | -104.7001
12.396 | -104.7067
12.3997 | -104.7105
12.4034 | -104.7149
12.4072 | -104.7192
12.4109 | -104.723
12.4146 | -104.7244
12.4184 | -104.7248
12.4221 | -104.7251
12.4258 | -104.7207
12.4295 | -104.7138
12.4333 | -104.7046
12.437 | -104.6938
12.4407 | -104.6842
12.4444 | -104.674
12.4482 | -104.6614
12.4519 | -104.651
12.4556 | -104.6434
12.4594 | -104.6361
12.4631 | -104.6286
12.4668 | -104.6194
12.4705 | -104.6122
12.4743 | -104.6093
12.478 | -104.6101
12.4817 | -104.6136
12.4854 | -104.6141
12.4892 | -104.6149
12.4929 | -104.6189
12.4966 | -104.6279
12.5003 | -104.6381
12.5041 | -104.6491
12.5078 | -104.6659
12.5115 | -104.6823
12.5153 | -104.6987
12.519 | -104.7151
12.5227 | -104.7331
12.5264 | -104.7497
12.5302 | -104.7662
12.5339 | -104.784
12.5376 | -104.8033
12.5413 | -104.8217
12.5451 | -104.8385
12.5488 | -104.8513
12.5525 | -104.8627
12.5563 | -104.8728
12.56 | -104.8814
12.5637 | -104.888
12.5674 | -104.8956
12.5712 | -104.9011
12.5749 | -104.9065
12.5786 | -104.9114
12.5823 | -104.9161
12.5861 | -104.9183
12.5898 | -104.9205
12.5935 | -104.9187
12.5973 | -104.9152
12.601 | -104.9131
12.6047 | -104.9104
12.6084 | -104.9093
12.6122 | -104.9087
12.6159 | -104.9069
12.6196 | -104.9048
12.6233 | -104.9009
12.6271 | -104.9009
12.6308 | -104.9031
12.6345 | -104.9042
12.6382 | -104.903
12.642 | -104.9035
12.6457 | -104.9042
12.6494 | -104.9053
12.6532 | -104.9066
12.6569 | -104.9036
12.6606 | -104.8958
12.6643 | -104.8878
12.6681 | -104.8826
12.6718 | -104.8773
12.6755 | -104.8696
12.6792 | -104.8597
12.683 | -104.8513
12.6867 | -104.8435
12.6904 | -104.8329
12.6942 | -104.8236
12.6979 | -104.8116
12.7016 | -104.8016
12.7053 | -104.7902
12.7091 | -104.781
12.7128 | -104.7726
12.7165 | -104.7637
12.7202 | -104.7555
12.724 | -104.7476
12.7277 | -104.7423
12.7314 | -104.7381
12.7352 | -104.7338
12.7389 | -104.7323
12.7426 | -104.7316
12.7463 | -104.7325
12.7501 | -104.7331
12.7538 | -104.7382
12.7575 | -104.746
12.7612 | -104.7568
12.765 | -104.7687
12.7687 | -104.7799
12.7724 | -104.7935
12.7761 | -104.8123
12.7799 | -104.8313
12.7836 | -104.8516
12.7873 | -104.8742
12.7911 | -104.8921
12.7948 | -104.9065
12.7985 | -104.9235
12.8022 | -104.9378
12.806 | -104.95
12.8097 | -104.9572
12.8134 | -104.961
12.8171 | -104.963
12.8209 | -104.9638
12.8246 | -104.9647
12.8283 | -104.9637
12.8321 | -104.9553
12.8358 | -104.9434
12.8395 | -104.9277
12.8432 | -104.9121
12.847 | -104.8942
12.8507 | -104.8732
12.8544 | -104.8524
12.8581 | -104.8288
12.8619 | -104.8057
12.8656 | -104.7842
12.8693 | -104.7605
12.873 | -104.7359
12.8768 | -104.711
12.8805 | -104.689
12.8842 | -104.6658
12.888 | -104.6478
12.8917 | -104.6314
12.8954 | -104.621
12.8991 | -104.6113
12.9029 | -104.6088
12.9066 | -104.6017
12.9103 | -104.5956
12.914 | -104.5903
12.9178 | -104.5891
12.9215 | -104.588
12.9252 | -104.5881
12.929 | -104.5908
12.9327 | -104.5904
12.9364 | -104.592
12.9401 | -104.5955
12.9439 | -104.5987
12.9476 | -104.6047
12.9513 | -104.6107
12.955 | -104.6193
12.9588 | -104.6295
12.9625 | -104.6391
12.9662 | -104.6518
12.97 | -104.663
12.9737 | -104.6755
12.9774 | -104.6938
12.9811 | -104.7154
12.9849 | -104.7315
12.9886 | -104.7486
12.9923 | -104.7678
12.996 | -104.7876
12.9998 | -104.8025
13.0035 | -104.8162
13.0072 | -104.8314
13.0109 | -104.8433
13.0147 | -104.8537
13.0184 | -104.8663
13.0221 | -104.8805
13.0259 | -104.8948
13.0296 | -104.9055
13.0333 | -104.9135
13.037 | -104.9204
13.0408 | -104.927
13.0445 | -104.9359
13.0482 | -104.9415
13.0519 | -104.9488
13.0557 | -104.9602
13.0594 | -104.9718
13.0631 | -104.9774
13.0669 | -104.9807
13.0706 | -104.9856
13.0743 | -104.9885
13.078 | -104.9878
13.0818 | -104.9783
13.0855 | -104.9684
13.0892 | -104.9607
13.0929 | -104.9505
13.0967 | -104.9383
13.1004 | -104.9245
13.1041 | -104.9103
13.1078 | -104.8901
13.1116 | -104.8684
13.1153 | -104.8455
13.119 | -104.8179
13.1228 | -104.793
13.1265 | -104.7663
13.1302 | -104.7391
13.1339 | -104.7118
13.1377 | -104.6814
13.1414 | -104.6516
13.1451 | -104.626
13.1488 | -104.5999
13.1526 | -104.5741
13.1563 | -104.5482
13.16 | -104.527
13.1638 | -104.5052
13.1675 | -104.4858
13.1712 | -104.4725
13.1749 | -104.4645
13.1787 | -104.4586
13.1824 | -104.4529
13.1861 | -104.4508
13.1898 | -104.457
13.1936 | -104.4667
13.1973 | -104.4759
13.201 | -104.4873
13.2048 | -104.5046
13.2085 | -104.5252
13.2122 | -104.5454
13.2159 | -104.566
13.2197 | -104.5922
13.2234 | -104.6218
13.2271 | -104.6499
13.2308 | -104.6795
13.2346 | -104.7102
13.2383 | -104.7416
13.242 | -104.7716
13.2457 | -104.7996
13.2495 | -104.8296
13.2532 | -104.8649
13.2569 | -104.9
13.2607 | -104.9295
13.2644 | -104.9569
13.2681 | -104.9859
13.2718 | -105.0191
13.2756 | -105.0501
13.2793 | -105.0742
13.283 | -105.0977
13.2867 | -105.1196
13.2905 | -105.1398
13.2942 | -105.1539
13.2979 | -105.1692
13.3017 | -105.187
13.3054 | -105.2022
13.3091 | -105.2173
13.3128 | -105.2294
13.3166 | -105.2387
13.3203 | -105.2476
13.324 | -105.2554
13.3277 | -105.263
13.3315 | -105.2697
13.3352 | -105.2729
13.3389 | -105.2752
13.3427 | -105.2788
13.3464 | -105.2784
13.3501 | -105.2769
13.3538 | -105.2786
13.3576 | -105.2771
13.3613 | -105.2701
13.365 | -105.2682
13.3687 | -105.2694
13.3725 | -105.2699
13.3762 | -105.2679
13.3799 | -105.2629
13.3836 | -105.2557
13.3874 | -105.2492
13.3911 | -105.2436
13.3948 | -105.2359
13.3986 | -105.2273
13.4023 | -105.2219
13.406 | -105.215
13.4097 | -105.204
13.4135 | -105.196
13.4172 | -105.188
13.4209 | -105.1754
13.4246 | -105.161
13.4284 | -105.1493
13.4321 | -105.1362
13.4358 | -105.1274
13.4396 | -105.1204
13.4433 | -105.1109
13.447 | -105.1035
13.4507 | -105.097
13.4545 | -105.0911
13.4582 | -105.0882
13.4619 | -105.081
13.4656 | -105.0717
13.4694 | -105.0641
13.4731 | -105.0644
13.4768 | -105.0647
13.4805 | -105.0655
13.4843 | -105.0685
13.488 | -105.0752
13.4917 | -105.0787
13.4955 | -105.0783
13.4992 | -105.0742
13.5029 | -105.0715
13.5066 | -105.0691
13.5104 | -105.0658
13.5141 | -105.0615
13.5178 | -105.0565
13.5215 | -105.0507
13.5253 | -105.0412
13.529 | -105.0276
13.5327 | -105.0127
13.5365 | -104.9972
13.5402 | -104.9827
13.5439 | -104.9655
13.5476 | -104.9451
13.5514 | -104.9261
13.5551 | -104.9071
13.5588 | -104.8852
13.5625 | -104.8632
13.5663 | -104.8415
13.57 | -104.8212
13.5737 | -104.8016
13.5775 | -104.786
13.5812 | -104.7731
13.5849 | -104.7611
13.5886 | -104.748
13.5924 | -104.7364
13.5961 | -104.7247
13.5998 | -104.7183
13.6035 | -104.7144
13.6073 | -104.71
13.611 | -104.7077
13.6147 | -104.7084
13.6184 | -104.712
13.6222 | -104.715
13.6259 | -104.7189
13.6296 | -104.7223
13.6334 | -104.7274
13.6371 | -104.7348
13.6408 | -104.7425
13.6445 | -104.7503
13.6483 | -104.76
13.652 | -104.7718
13.6557 | -104.7834
13.6594 | -104.7982
13.6632 | -104.8147
13.6669 | -104.8297
13.6706 | -104.8385
13.6744 | -104.8472
13.6781 | -104.8588
13.6818 | -104.8714
13.6855 | -104.8823
13.6893 | -104.8943
13.693 | -104.9062
13.6967 | -104.921
13.7004 | -104.9357
13.7042 | -104.9495
13.7079 | -104.9649
13.7116 | -104.9805
13.7154 | -104.9982
13.7191 | -105.018
13.7228 | -105.0396
13.7265 | -105.0609
13.7303 | -105.0792
13.734 | -105.0978
13.7377 | -105.1194
13.7414 | -105.1369
13.7452 | -105.1531
13.7489 | -105.1653
13.7526 | -105.1771
13.7563 | -105.1893
13.7601 | -105.197
13.7638 | -105.2066
13.7675 | -105.2148
13.7713 | -105.2192
13.775 | -105.2208
13.7787 | -105.2218
13.7824 | -105.221
13.7862 | -105.2187
13.7899 | -105.2152
13.7936 | -105.2141
13.7973 | -105.2131
13.8011 | -105.2129
13.8048 | -105.2118
13.8085 | -105.2126
13.8123 | -105.2098
13.816 | -105.2067
13.8197 | -105.2091
13.8234 | -105.2142
13.8272 | -105.2168
13.8309 | -105.2195
13.8346 | -105.2238
13.8383 | -105.2281
13.8421 | -105.2314
13.8458 | -105.2335
13.8495 | -105.2368
13.8532 | -105.2405
13.857 | -105.2468
13.8607 | -105.2501
13.8644 | -105.2526
13.8682 | -105.256
13.8719 | -105.2584
13.8756 | -105.2556
13.8793 | -105.2516
13.8831 | -105.2509
13.8868 | -105.2489
13.8905 | -105.247
13.8942 | -105.2412
13.898 | -105.2339
13.9017 | -105.2244
13.9054 | -105.2142
13.9092 | -105.2053
13.9129 | -105.1918
13.9166 | -105.1768
13.9203 | -105.1633
13.9241 | -105.1473
13.9278 | -105.1326
13.9315 | -105.1176
13.9352 | -105.1015
13.939 | -105.0914
13.9427 | -105.0805
13.9464 | -105.0698
13.9502 | -105.0588
13.9539 | -105.0439
13.9576 | -105.0327
13.9613 | -105.0131
13.9651 | -104.9858
13.9688 | -104.964
13.9725 | -104.9478
13.9762 | -104.9309
13.98 | -104.9156
13.9837 | -104.898
13.9874 | -104.8806
13.9911 | -104.8627
13.9949 | -104.8394
13.9986 | -104.8132
14.0023 | -104.789
14.0061 | -104.7658
14.0098 | -104.7409
14.0135 | -104.7176
14.0172 | -104.6916
14.021 | -104.661
14.0247 | -104.6302
14.0284 | -104.6028
14.0321 | -104.5764
14.0359 | -104.5504
14.0396 | -104.5279
14.0433 | -104.5079
14.0471 | -104.4914
14.0508 | -104.4769
14.0545 | -104.4645
14.0582 | -104.4564
14.062 | -104.4554
14.0657 | -104.4601
14.0694 | -104.4672
14.0731 | -104.4807
14.0769 | -104.5012
14.0806 | -104.5235
14.0843 | -104.5474
14.0881 | -104.5774
14.0918 | -104.6107
14.0955 | -104.6468
14.0992 | -104.6825
14.103 | -104.7216
14.1067 | -104.7614
14.1104 | -104.7981
14.1141 | -104.8389
14.1179 | -104.8807
14.1216 | -104.918
14.1253 | -104.9564
14.129 | -104.9889
14.1328 | -105.0124
14.1365 | -105.0363
14.1402 | -105.0561
14.144 | -105.0748
14.1477 | -105.0933
14.1514 | -105.1087
14.1551 | -105.1221
14.1589 | -105.1303
14.1626 | -105.1387
14.1663 | -105.1508
14.17 | -105.154
14.1738 | -105.1475
14.1775 | -105.1364
14.1812 | -105.1231
14.185 | -105.1082
14.1887 | -105.0939
14.1924 | -105.08
14.1961 | -105.0675
14.1999 | -105.0564
14.2036 | -105.0462
14.2073 | -105.0344
14.211 | -105.0218
14.2148 | -105.0128
14.2185 | -105.0079
14.2222 | -105.0056
14.2259 | -105.0045
14.2297 | -105.0031
14.2334 | -105.0022
14.2371 | -105.0051
14.2409 | -105.012
14.2446 | -105.0213
14.2483 | -105.032
14.252 | -105.0457
14.2558 | -105.0628
14.2595 | -105.0796
14.2632 | -105.0962
14.2669 | -105.1131
14.2707 | -105.1308
14.2744 | -105.1489
14.2781 | -105.1611
14.2819 | -105.1728
14.2856 | -105.1878
14.2893 | -105.2019
14.293 | -105.213
14.2968 | -105.2241
14.3005 | -105.2344
14.3042 | -105.2454
14.3079 | -105.2512
14.3117 | -105.2543
14.3154 | -105.2578
14.3191 | -105.2591
14.3229 | -105.2569
14.3266 | -105.253
14.3303 | -105.2459
14.334 | -105.2425
14.3378 | -105.2417
14.3415 | -105.2359
14.3452 | -105.2291
14.3489 | -105.2238
14.3527 | -105.2182
14.3564 | -105.2136
14.3601 | -105.2094
14.3638 | -105.2066
14.3676 | -105.207
14.3713 | -105.2068
14.375 | -105.2081
14.3788 | -105.2097
14.3825 | -105.2116
14.3862 | -105.2149
14.3899 | -105.2229
14.3937 | -105.2306
14.3974 | -105.2384
14.4011 | -105.2464
14.4048 | -105.2557
14.4086 | -105.2637
14.4123 | -105.2708
14.416 | -105.2813
14.4198 | -105.2892
14.4235 | -105.293
14.4272 | -105.2984
14.4309 | -105.3054
14.4347 | -105.3131
14.4384 | -105.3191
14.4421 | -105.3174
14.4458 | -105.3139
14.4496 | -105.3112
14.4533 | -105.309
14.457 | -105.305
14.4608 | -105.3005
14.4645 | -105.2975
14.4682 | -105.2929
14.4719 | -105.2877
14.4757 | -105.2821
14.4794 | -105.273
14.4831 | -105.2667
14.4868 | -105.2609
14.4906 | -105.2545
14.4943 | -105.2494
14.498 | -105.2437
14.5017 | -105.2363
14.5055 | -105.2293
14.5092 | -105.2221
14.5129 | -105.2186
14.5167 | -105.2206
14.5204 | -105.2204
14.5241 | -105.2176
14.5278 | -105.2145
14.5316 | -105.2138
14.5353 | -105.2152
14.539 | -105.2156
14.5427 | -105.2149
14.5465 | -105.2126
14.5502 | -105.2093
14.5539 | -105.2056
14.5577 | -105.2003
14.5614 | -105.1965
14.5651 | -105.195
14.5688 | -105.1932
14.5726 | -105.1883
14.5763 | -105.1816
14.58 | -105.1718
14.5837 | -105.1631
14.5875 | -105.1546
14.5912 | -105.1472
14.5949 | -105.1394
14.5986 | -105.1279
14.6024 | -105.1174
14.6061 | -105.1042
14.6098 | -105.0902
14.6136 | -105.0752
14.6173 | -105.0624
14.621 | -105.0519
14.6247 | -105.038
14.6285 | -105.024
14.6322 | -105.0109
14.6359 | -104.9979
14.6396 | -104.9863
14.6434 | -104.9758
14.6471 | -104.9663
14.6508 | -104.9623
14.6546 | -104.9588
14.6583 | -104.9545
14.662 | -104.9507
14.6657 | -104.9491
14.6695 | -104.9495
14.6732 | -104.9516
14.6769 | -104.9577
14.6806 | -104.9667
14.6844 | -104.9755
14.6881 | -104.9887
14.6918 | -104.9997
14.6956 | -105.0095
14.6993 | -105.0222
14.703 | -105.0381
14.7067 | -105.0561
14.7105 | -105.0751
14.7142 | -105.094
14.7179 | -105.1114
14.7216 | -105.1292
14.7254 | -105.1454
14.7291 | -105.1616
14.7328 | -105.1795
14.7365 | -105.1977
14.7403 | -105.217
14.744 | -105.2339
14.7477 | -105.2512
14.7515 | -105.2693
14.7552 | -105.2842
14.7589 | -105.3006
14.7626 | -105.3168
14.7664 | -105.333
14.7701 | -105.3476
14.7738 | -105.3579
14.7775 | -105.3658
14.7813 | -105.3751
14.785 | -105.3856
14.7887 | -105.3975
14.7925 | -105.4109
14.7962 | -105.4182
14.7999 | -105.4235
14.8036 | -105.43
14.8074 | -105.4364
14.8111 | -105.441
14.8148 | -105.4455
14.8185 | -105.4517
14.8223 | -105.4589
14.826 | -105.463
14.8297 | -105.4658
14.8334 | -105.4694
14.8372 | -105.4757
14.8409 | -105.4812
14.8446 | -105.4869
14.8484 | -105.4889
14.8521 | -105.4894
14.8558 | -105.4931
14.8595 | -105.4985
14.8633 | -105.5018
14.867 | -105.5012
14.8707 | -105.4981
14.8744 | -105.4939
14.8782 | -105.4877
14.8819 | -105.4816
14.8856 | -105.4754
14.8894 | -105.4673
14.8931 | -105.4546
14.8968 | -105.4398
14.9005 | -105.4252
14.9043 | -105.4122
14.908 | -105.3957
14.9117 | -105.3738
14.9154 | -105.3505
14.9192 | -105.3293
14.9229 | -105.3095
14.9266 | -105.2918
14.9304 | -105.2778
14.9341 | -105.2634
14.9378 | -105.252
14.9415 | -105.2391
14.9453 | -105.2269
14.949 | -105.2176
14.9527 | -105.2092
14.9564 | -105.2054
14.9602 | -105.2035
14.9639 | -105.2015
14.9676 | -105.2018
14.9713 | -105.2044
14.9751 | -105.2104
14.9788 | -105.2156
14.9825 | -105.2242
14.9863 | -105.2347
14.99 | -105.246
14.9937 | -105.2563
14.9974 | -105.2677
15.0012 | -105.2749
15.0049 | -105.2832
15.0086 | -105.294
15.0123 | -105.3038
15.0161 | -105.3093
15.0198 | -105.3143
15.0235 | -105.3204
15.0273 | -105.3271
15.031 | -105.3311
15.0347 | -105.3311
15.0384 | -105.3326
15.0422 | -105.3329
15.0459 | -105.3297
15.0496 | -105.3244
15.0533 | -105.3137
15.0571 | -105.3093
15.0608 | -105.3025
15.0645 | -105.2984
15.0683 | -105.2928
15.072 | -105.2841
15.0757 | -105.2766
15.0794 | -105.2699
15.0832 | -105.264
15.0869 | -105.2581
15.0906 | -105.2535
15.0943 | -105.2493
15.0981 | -105.2399
15.1018 | -105.2318
15.1055 | -105.2249
15.1092 | -105.2201
15.113 | -105.2162
15.1167 | -105.2117
15.1204 | -105.2063
15.1242 | -105.2027
15.1279 | -105.1989
15.1316 | -105.1906
15.1353 | -105.1812
15.1391 | -105.1734
15.1428 | -105.1674
15.1465 | -105.159
15.1502 | -105.15
15.154 | -105.1425
15.1577 | -105.1372
15.1614 | -105.1314
15.1652 | -105.1259
15.1689 | -105.1196
15.1726 | -105.1131
15.1763 | -105.1077
15.1801 | -105.1041
15.1838 | -105.1021
15.1875 | -105.0999
15.1912 | -105.0969
15.195 | -105.0984
15.1987 | -105.0975
15.2024 | -105.0962
15.2061 | -105.0975
15.2099 | -105.1015
15.2136 | -105.1049
15.2173 | -105.1147
15.2211 | -105.1222
15.2248 | -105.1284
15.2285 | -105.1402
15.2322 | -105.1524
15.236 | -105.1655
15.2397 | -105.1755
15.2434 | -105.1857
15.2471 | -105.1966
15.2509 | -105.208
15.2546 | -105.2172
15.2583 | -105.227
15.2621 | -105.2369
15.2658 | -105.2457
15.2695 | -105.2538
15.2732 | -105.2591
15.277 | -105.2656
15.2807 | -105.2722
15.2844 | -105.2739
15.2881 | -105.2751
15.2919 | -105.2766
15.2956 | -105.2777
15.2993 | -105.2808
15.3031 | -105.2829
15.3068 | -105.2847
15.3105 | -105.286
15.3142 | -105.2887
15.318 | -105.2909
15.3217 | -105.2904
15.3254 | -105.2906
15.3291 | -105.2923
15.3329 | -105.2952
15.3366 | -105.2949
15.3403 | -105.2956
15.344 | -105.2949
15.3478 | -105.297
15.3515 | -105.3032
15.3552 | -105.3066
15.359 | -105.3065
15.3627 | -105.3058
15.3664 | -105.3062
15.3701 | -105.3054
15.3739 | -105.3026
15.3776 | -105.2998
15.3813 | -105.2969
15.385 | -105.2962
15.3888 | -105.2967
15.3925 | -105.2947
15.3962 | -105.2915
15.4 | -105.2908
15.4037 | -105.2889
15.4074 | -105.286
15.4111 | -105.285
15.4149 | -105.2824
15.4186 | -105.2805
15.4223 | -105.2786
15.426 | -105.2713
15.4298 | -105.2616
15.4335 | -105.2559
15.4372 | -105.2481
15.441 | -105.2407
15.4447 | -105.2359
15.4484 | -105.2301
15.4521 | -105.2257
15.4559 | -105.2187
15.4596 | -105.2155
15.4633 | -105.2118
15.467 | -105.2075
15.4708 | -105.2044
15.4745 | -105.2038
15.4782 | -105.2013
15.4819 | -105.201
15.4857 | -105.2002
15.4894 | -105.196
15.4931 | -105.1945
15.4969 | -105.1903
15.5006 | -105.1877
15.5043 | -105.1885
15.508 | -105.1894
15.5118 | -105.1886
15.5155 | -105.1845
15.5192 | -105.1763
15.5229 | -105.1691
15.5267 | -105.1646
15.5304 | -105.1607
15.5341 | -105.1563
15.5379 | -105.1521
15.5416 | -105.1443
15.5453 | -105.1337
15.549 | -105.1195
15.5528 | -105.1063
15.5565 | -105.0961
15.5602 | -105.0818
15.5639 | -105.0694
15.5677 | -105.0544
15.5714 | -105.0357
15.5751 | -105.0197
15.5788 | -105.0071
15.5826 | -104.9901
15.5863 | -104.9747
15.59 | -104.9629
15.5938 | -104.9533
15.5975 | -104.9423
15.6012 | -104.9314
15.6049 | -104.924
15.6087 | -104.9153
15.6124 | -104.9065
15.6161 | -104.9024
15.6198 | -104.9007
15.6236 | -104.8977
15.6273 | -104.8983
15.631 | -104.9014
15.6348 | -104.9054
15.6385 | -104.9109
15.6422 | -104.9163
15.6459 | -104.924
15.6497 | -104.9358
15.6534 | -104.9481
15.6571 | -104.958
15.6608 | -104.9647
15.6646 | -104.973
15.6683 | -104.9824
15.672 | -104.9878
15.6758 | -104.9924
15.6795 | -105.0021
15.6832 | -105.0104
15.6869 | -105.0153
15.6907 | -105.022
15.6944 | -105.026
15.6981 | -105.0283
15.7018 | -105.0302
15.7056 | -105.028
15.7093 | -105.026
15.713 | -105.0261
15.7167 | -105.0256
15.7205 | -105.0244
15.7242 | -105.0208
15.7279 | -105.0162
15.7317 | -105.0103
15.7354 | -105.0035
15.7391 | -105
15.7428 | -104.9983
15.7466 | -104.9971
15.7503 | -104.9941
15.754 | -104.9893
15.7577 | -104.9829
15.7615 | -104.9772
15.7652 | -104.9739
15.7689 | -104.9715
15.7727 | -104.9718
15.7764 | -104.9715
15.7801 | -104.9696
15.7838 | -104.9713
15.7876 | -104.9726
15.7913 | -104.9754
15.795 | -104.9824
15.7987 | -104.9883
15.8025 | -104.9945
15.8062 | -105.0028
15.8099 | -105.0133
15.8137 | -105.0272
15.8174 | -105.0393
15.8211 | -105.0497
15.8248 | -105.0634
15.8286 | -105.0786
15.8323 | -105.0937
15.836 | -105.1095
15.8397 | -105.1275
15.8435 | -105.146
15.8472 | -105.1637
15.8509 | -105.1819
15.8546 | -105.1968
15.8584 | -105.2109
15.8621 | -105.2237
15.8658 | -105.2356
15.8696 | -105.2491
15.8733 | -105.2576
15.877 | -105.2625
15.8807 | -105.2664
15.8845 | -105.2685
15.8882 | -105.2721
15.8919 | -105.2738
15.8956 | -105.2723
15.8994 | -105.2721
15.9031 | -105.2717
15.9068 | -105.2661
15.9106 | -105.2597
15.9143 | -105.2553
15.918 | -105.2503
15.9217 | -105.2475
15.9255 | -105.2458
15.9292 | -105.2434
15.9329 | -105.2433
15.9366 | -105.2437
15.9404 | -105.244
15.9441 | -105.2425
15.9478 | -105.2433
15.9515 | -105.2446
15.9553 | -105.2452
15.959 | -105.248
15.9627 | -105.2552
15.9665 | -105.2592
15.9702 | -105.2626
15.9739 | -105.2676
15.9776 | -105.272
15.9814 | -105.277
15.9851 | -105.2841
15.9888 | -105.2965
15.9925 | -105.3088
15.9963 | -105.3203
16 | -105.3317
16.0037 | -105.342
16.0075 | -105.353
16.0112 | -105.3643
16.0149 | -105.3751
16.0186 | -105.3855
16.0224 | -105.3974
16.0261 | -105.411
16.0298 | -105.4218
16.0335 | -105.4354
16.0373 | -105.4466
16.041 | -105.4549
16.0447 | -105.4651
16.0485 | -105.4747
16.0522 | -105.4833
16.0559 | -105.4899
16.0596 | -105.4939
16.0634 | -105.4975
16.0671 | -105.4981
16.0708 | -105.4953
16.0745 | -105.4915
16.0783 | -105.4882
16.082 | -105.4863
16.0857 | -105.4814
16.0894 | -105.4728
16.0932 | -105.4635
16.0969 | -105.4536
16.1006 | -105.4448
16.1044 | -105.4341
16.1081 | -105.4197
16.1118 | -105.4025
16.1155 | -105.3836
16.1193 | -105.364
16.123 | -105.3437
16.1267 | -105.3232
16.1304 | -105.2999
16.1342 | -105.2781
16.1379 | -105.2555
16.1416 | -105.2352
16.1454 | -105.2163
16.1491 | -105.195
16.1528 | -105.1739
16.1565 | -105.1521
16.1603 | -105.1307
16.164 | -105.1107
16.1677 | -105.0942
16.1714 | -105.078
16.1752 | -105.0622
16.1789 | -105.0458
16.1826 | -105.0324
16.1864 | -105.0222
16.1901 | -105.0101
16.1938 | -104.9968
16.1975 | -104.9891
16.2013 | -104.9842
16.205 | -104.9804
16.2087 | -104.9806
16.2124 | -104.9816
16.2162 | -104.9856
16.2199 | -104.9906
16.2236 | -104.9988
16.2273 | -105.0087
16.2311 | -105.0186
16.2348 | -105.0305
16.2385 | -105.0465
16.2423 | -105.0644
16.246 | -105.0842
16.2497 | -105.1037
16.2534 | -105.122
16.2572 | -105.1408
16.2609 | -105.1622
16.2646 | -105.1836
16.2683 | -105.2016
16.2721 | -105.2195
16.2758 | -105.2349
16.2795 | -105.2516
16.2833 | -105.2691
16.287 | -105.2863
16.2907 | -105.3032
16.2944 | -105.3175
16.2982 | -105.33
16.3019 | -105.3427
16.3056 | -105.3553
16.3093 | -105.3635
16.3131 | -105.3713
16.3168 | -105.3784
16.3205 | -105.3852
16.3242 | -105.3906
16.328 | -105.3939
16.3317 | -105.3952
16.3354 | -105.3982
16.3392 | -105.4035
16.3429 | -105.4072
16.3466 | -105.4096
16.3503 | -105.41
16.3541 | -105.4097
16.3578 | -105.4068
16.3615 | -105.4033
16.3652 | -105.3999
16.369 | -105.3971
16.3727 | -105.3976
16.3764 | -105.394
16.3802 | -105.3888
16.3839 | -105.3858
16.3876 | -105.3827
16.3913 | -105.3812
16.3951 | -105.3796
16.3988 | -105.3755
16.4025 | -105.3691
16.4062 | -105.3639
16.41 | -105.3635
16.4137 | -105.364
16.4174 | -105.365
16.4212 | -105.3672
16.4249 | -105.3722
16.4286 | -105.3774
16.4323 | -105.3788
16.4361 | -105.3792
16.4398 | -105.3795
16.4435 | -105.3822
16.4472 | -105.387
16.451 | -105.3908
16.4547 | -105.3925
16.4584 | -105.3939
16.4621 | -105.398
16.4659 | -105.4041
16.4696 | -105.4104
16.4733 | -105.4144
16.4771 | -105.417
16.4808 | -105.4181
16.4845 | -105.4204
16.4882 | -105.4223
16.492 | -105.427
16.4957 | -105.4324
16.4994 | -105.4371
16.5031 | -105.4442
16.5069 | -105.4521
16.5106 | -105.459
16.5143 | -105.4614
16.5181 | -105.4627
16.5218 | -105.4644
16.5255 | -105.4645
16.5292 | -105.4683
16.533 | -105.4725
16.5367 | -105.4754
16.5404 | -105.479
16.5441 | -105.4854
16.5479 | -105.4887
16.5516 | -105.4919
16.5553 | -105.4968
16.559 | -105.5022
16.5628 | -105.5053
16.5665 | -105.5092
16.5702 | -105.5131
16.574 | -105.5166
16.5777 | -105.5214
16.5814 | -105.5258
16.5851 | -105.531
16.5889 | -105.5367
16.5926 | -105.5442
16.5963 | -105.5499
16.6 | -105.5565
16.6038 | -105.5639
16.6075 | -105.5669
16.6112 | -105.5754
16.615 | -105.5851
16.6187 | -105.5935
16.6224 | -105.6018
16.6261 | -105.6102
16.6299 | -105.6229
16.6336 | -105.6364
16.6373 | -105.6523
16.641 | -105.6677
16.6448 | -105.6805
16.6485 | -105.6949
16.6522 | -105.7094
16.656 | -105.7239
16.6597 | -105.7396
16.6634 | -105.7552
16.6671 | -105.7693
16.6709 | -105.7853
16.6746 | -105.8015
16.6783 | -105.8153
16.682 | -105.8251
16.6858 | -105.8358
16.6895 | -105.8447
16.6932 | -105.8546
16.6969 | -105.8648
16.7007 | -105.8718
16.7044 | -105.8763
16.7081 | -105.8781
16.7119 | -105.8786
16.7156 | -105.8808
16.7193 | -105.8825
16.723 | -105.8807
16.7268 | -105.8749
16.7305 | -105.8614
16.7342 | -105.8493
16.7379 | -105.8375
16.7417 | -105.8247
16.7454 | -105.8116
16.7491 | -105.8004
16.7529 | -105.7874
16.7566 | -105.7707
16.7603 | -105.7523
16.764 | -105.7384
16.7678 | -105.7241
16.7715 | -105.7088
16.7752 | -105.6924
16.7789 | -105.6768
16.7827 | -105.6621
16.7864 | -105.6502
16.7901 | -105.6387
16.7938 | -105.6251
16.7976 | -105.6113
16.8013 | -105.6001
16.805 | -105.5904
16.8088 | -105.5811
16.8125 | -105.5729
16.8162 | -105.5665
16.8199 | -105.5629
16.8237 | -105.5594
16.8274 | -105.5565
16.8311 | -105.5545
16.8348 | -105.5546
16.8386 | -105.5533
16.8423 | -105.5533
16.846 | -105.5563
16.8498 | -105.5559
16.8535 | -105.5559
16.8572 | -105.5573
16.8609 | -105.5589
16.8647 | -105.5616
16.8684 | -105.5645
16.8721 | -105.5676
16.8758 | -105.5689
16.8796 | -105.5725
16.8833 | -105.5746
16.887 | -105.5783
16.8908 | -105.5815
16.8945 | -105.5831
16.8982 | -105.5859
16.9019 | -105.5904
16.9057 | -105.5946
16.9094 | -105.5976
16.9131 | -105.6019
16.9168 | -105.6057
16.9206 | -105.6109
16.9243 | -105.6158
16.928 | -105.6183
16.9317 | -105.6209
16.9355 | -105.6258
16.9392 | -105.6294
16.9429 | -105.6328
16.9467 | -105.6366
16.9504 | -105.6404
16.9541 | -105.6447
16.9578 | -105.6458
16.9616 | -105.645
16.9653 | -105.6473
16.969 | -105.6497
16.9727 | -105.6518
16.9765 | -105.654
16.9802 | -105.6566
16.9839 | -105.6588
16.9877 | -105.6592
16.9914 | -105.658
16.9951 | -105.6602
16.9988 | -105.6656
17.0026 | -105.6693
17.0063 | -105.6706
17.01 | -105.6702
17.0137 | -105.669
17.0175 | -105.6669
17.0212 | -105.6655
17.0249 | -105.6622
17.0287 | -105.6597
17.0324 | -105.6603
17.0361 | -105.6609
17.0398 | -105.6618
17.0436 | -105.6619
17.0473 | -105.6624
17.051 | -105.6643
17.0547 | -105.6657
17.0585 | -105.6671
17.0622 | -105.6673
17.0659 | -105.6691
17.0696 | -105.6703
17.0734 | -105.6733
17.0771 | -105.6785
17.0808 | -105.6854
17.0846 | -105.6938
17.0883 | -105.7017
17.092 | -105.7104
17.0957 | -105.7211
17.0995 | -105.7326
17.1032 | -105.7432
17.1069 | -105.7534
17.1106 | -105.7644
17.1144 | -105.777
17.1181 | -105.7902
17.1218 | -105.8013
17.1256 | -105.813
17.1293 | -105.8238
17.133 | -105.8363
17.1367 | -105.8497
17.1405 | -105.8612
17.1442 | -105.8701
17.1479 | -105.8806
17.1516 | -105.89
17.1554 | -105.8978
17.1591 | -105.9066
17.1628 | -105.9144
17.1666 | -105.9232
17.1703 | -105.9293
17.174 | -105.9349
17.1777 | -105.9386
17.1815 | -105.9434
17.1852 | -105.9464
17.1889 | -105.9523
17.1926 | -105.9589
17.1964 | -105.9603
17.2001 | -105.9623
17.2038 | -105.963
17.2075 | -105.9672
17.2113 | -105.9683
17.215 | -105.9716
17.2187 | -105.9754
17.2225 | -105.9738
17.2262 | -105.9712
17.2299 | -105.97
17.2336 | -105.9677
17.2374 | -105.9659
17.2411 | -105.9655
17.2448 | -105.9636
17.2485 | -105.9621
17.2523 | -105.9591
17.256 | -105.957
17.2597 | -105.9554
17.2635 | -105.9579
17.2672 | -105.9584
17.2709 | -105.959
17.2746 | -105.9565
17.2784 | -105.9539
17.2821 | -105.9543
17.2858 | -105.9544
17.2895 | -105.955
17.2933 | -105.9521
17.297 | -105.9477
17.3007 | -105.948
17.3044 | -105.9477
17.3082 | -105.9493
17.3119 | -105.9496
17.3156 | -105.9493
17.3194 | -105.9494
17.3231 | -105.9483
17.3268 | -105.9468
17.3305 | -105.9431
17.3343 | -105.9398
17.338 | -105.9359
17.3417 | -105.9319
17.3454 | -105.928
17.3492 | -105.9267
17.3529 | -105.9224
17.3566 | -105.9154
17.3604 | -105.9095
17.3641 | -105.9018
17.3678 | -105.8935
17.3715 | -105.8851
17.3753 | -105.8758
17.379 | -105.8644
17.3827 | -105.8541
17.3864 | -105.8449
17.3902 | -105.8352
17.3939 | -105.8248
17.3976 | -105.8144
17.4014 | -105.8039
17.4051 | -105.7933
17.4088 | -105.7831
17.4125 | -105.772
17.4163 | -105.7643
17.42 | -105.7564
17.4237 | -105.7453
17.4274 | -105.7347
17.4312 | -105.7266
17.4349 | -105.7173
17.4386 | -105.7083
17.4423 | -105.7041
17.4461 | -105.7007
17.4498 | -105.6974
17.4535 | -105.6921
17.4573 | -105.6899
17.461 | -105.6867
17.4647 | -105.6818
17.4684 | -105.6802
17.4722 | -105.6807
17.4759 | -105.6813
17.4796 | -105.6799
17.4833 | -105.6794
17.4871 | -105.6797
17.4908 | -105.6775
17.4945 | -105.6768
17.4983 | -105.6781
17.502 | -105.6809
17.5057 | -105.6832
17.5094 | -105.6872
17.5132 | -105.69
17.5169 | -105.6901
17.5206 | -105.6893
17.5243 | -105.6895
17.5281 | -105.6896
17.5318 | -105.6893
17.5355 | -105.6876
17.5392 | -105.6854
17.543 | -105.6861
17.5467 | -105.6866
17.5504 | -105.6837
17.5542 | -105.682
17.5579 | -105.6789
17.5616 | -105.6763
17.5653 | -105.6762
17.5691 | -105.6715
17.5728 | -105.6662
17.5765 | -105.6636
17.5802 | -105.664
17.584 | -105.6658
17.5877 | -105.6671
17.5914 | -105.6675
17.5952 | -105.6679
17.5989 | -105.6677
17.6026 | -105.6695
17.6063 | -105.6701
17.6101 | -105.6711
17.6138 | -105.6738
17.6175 | -105.6791
17.6212 | -105.6846
17.625 | -105.69
17.6287 | -105.6949
17.6324 | -105.7004
17.6362 | -105.7056
17.6399 | -105.7135
17.6436 | -105.7192
17.6473 | -105.7256
17.6511 | -105.7327
17.6548 | -105.7369
17.6585 | -105.7415
17.6622 | -105.7456
17.666 | -105.7516
17.6697 | -105.758
17.6734 | -105.7668
17.6771 | -105.7748
17.6809 | -105.7801
17.6846 | -105.7846
17.6883 | -105.789
17.6921 | -105.7928
17.6958 | -105.7947
17.6995 | -105.7946
17.7032 | -105.7973
17.707 | -105.7967
17.7107 | -105.796
17.7144 | -105.7961
17.7181 | -105.7933
17.7219 | -105.7886
17.7256 | -105.7821
17.7293 | -105.7774
17.7331 | -105.7733
17.7368 | -105.767
17.7405 | -105.752
17.7442 | -105.7449
17.748 | -105.7364
17.7517 | -105.728
17.7554 | -105.7188
17.7591 | -105.7101
17.7629 | -105.7013
17.7666 | -105.6917
17.7703 | -105.684
17.7741 | -105.6776
17.7778 | -105.671
17.7815 | -105.6643
17.7852 | -105.657
17.789 | -105.6501
17.7927 | -105.6459
17.7964 | -105.6421
17.8001 | -105.6388
17.8039 | -105.6372
17.8076 | -105.6369
17.8113 | -105.6372
17.815 | -105.6381
17.8188 | -105.6375
17.8225 | -105.6423
17.8262 | -105.649
17.83 | -105.6534
17.8337 | -105.656
17.8374 | -105.6579
17.8411 | -105.6617
17.8449 | -105.6663
17.8486 | -105.672
17.8523 | -105.6778
17.856 | -105.6837
17.8598 | -105.6906
17.8635 | -105.6941
17.8672 | -105.6987
17.871 | -105.7014
17.8747 | -105.7042
17.8784 | -105.7075
17.8821 | -105.7115
17.8859 | -105.7122
17.8896 | -105.7127
17.8933 | -105.7147
17.897 | -105.7183
17.9008 | -105.7208
17.9045 | -105.7221
17.9082 | -105.7252
17.912 | -105.7276
17.9157 | -105.7314
17.9194 | -105.7311
17.9231 | -105.7311
17.9269 | -105.7325
17.9306 | -105.7351
17.9343 | -105.738
17.938 | -105.7409
17.9418 | -105.7427
17.9455 | -105.7435
17.9492 | -105.7472
17.9529 | -105.7531
17.9567 | -105.7574
17.9604 | -105.7635
17.9641 | -105.7702
17.9679 | -105.7768
17.9716 | -105.7816
17.9753 | -105.7866
17.979 | -105.7914
17.9828 | -105.7969
17.9865 | -105.8018
17.9902 | -105.8054
17.9939 | -105.8114
17.9977 | -105.816
18.0014 | -105.8176
18.0051 | -105.8196
18.0089 | -105.8205
18.0126 | -105.8223
18.0163 | -105.8247
18.02 | -105.825
18.0238 | -105.8245
18.0275 | -105.8244
18.0312 | -105.8249
18.0349 | -105.824
18.0387 | -105.8199
18.0424 | -105.8158
18.0461 | -105.815
18.0498 | -105.8117
18.0536 | -105.8069
18.0573 | -105.8019
18.061 | -105.7938
18.0648 | -105.7845
18.0685 | -105.7771
18.0722 | -105.7694
18.0759 | -105.7615
18.0797 | -105.7509
18.0834 | -105.7411
18.0871 | -105.7335
18.0908 | -105.7262
18.0946 | -105.7199
18.0983 | -105.7128
18.102 | -105.7059
18.1058 | -105.699
18.1095 | -105.6919
18.1132 | -105.6855
18.1169 | -105.6791
18.1207 | -105.6767
18.1244 | -105.6733
18.1281 | -105.6695
18.1318 | -105.6657
18.1356 | -105.6614
18.1393 | -105.6567
18.143 | -105.6569
18.1468 | -105.6572
18.1505 | -105.6556
18.1542 | -105.6543
18.1579 | -105.6542
18.1617 | -105.655
18.1654 | -105.6573
18.1691 | -105.6603
18.1728 | -105.6618
18.1766 | -105.6673
18.1803 | -105.6738
18.184 | -105.679
18.1877 | -105.6865
18.1915 | -105.6946
18.1952 | -105.7038
18.1989 | -105.714
18.2027 | -105.7239
18.2064 | -105.7333
18.2101 | -105.7413
18.2138 | -105.7502
18.2176 | -105.7606
18.2213 | -105.7709
18.225 | -105.78
18.2287 | -105.7907
18.2325 | -105.8001
18.2362 | -105.8117
18.2399 | -105.8211
18.2437 | -105.8296
18.2474 | -105.8374
18.2511 | -105.845
18.2548 | -105.8511
18.2586 | -105.8553
18.2623 | -105.8589
18.266 | -105.8631
18.2697 | -105.8665
18.2735 | -105.8691
18.2772 | -105.8719
18.2809 | -105.8699
18.2846 | -105.8741
18.2884 | -105.8764
18.2921 | -105.8757
18.2958 | -105.877
18.2996 | -105.8796
18.3033 | -105.8783
18.307 | -105.8743
18.3107 | -105.8718
18.3145 | -105.8684
18.3182 | -105.8646
18.3219 | -105.8624
18.3256 | -105.8594
18.3294 | -105.8525
18.3331 | -105.8464
18.3368 | -105.8416
18.3406 | -105.8357
18.3443 | -105.8284
18.348 | -105.8206
18.3517 | -105.8131
18.3555 | -105.8056
18.3592 | -105.7977
18.3629 | -105.7875
18.3666 | -105.778
18.3704 | -105.7706
18.3741 | -105.7638
18.3778 | -105.7566
18.3816 | -105.7466
18.3853 | -105.7364
18.389 | -105.7275
18.3927 | -105.7176
18.3965 | -105.7096
18.4002 | -105.7026
18.4039 | -105.6936
18.4076 | -105.6853
18.4114 | -105.6768
18.4151 | -105.67
18.4188 | -105.6641
18.4225 | -105.6582
18.4263 | -105.6517
18.43 | -105.6454
18.4337 | -105.6404
18.4375 | -105.6331
18.4412 | -105.6254
18.4449 | -105.62
18.4486 | -105.6148
18.4524 | -105.6105
18.4561 | -105.6068
18.4598 | -105.6048
18.4635 | -105.6036
18.4673 | -105.6028
18.471 | -105.6031
18.4747 | -105.6041
18.4785 | -105.6072
18.4822 | -105.6105
18.4859 | -105.6129
18.4896 | -105.6154
18.4934 | -105.6222
18.4971 | -105.6275
18.5008 | -105.6329
18.5045 | -105.6388
18.5083 | -105.6457
18.512 | -105.6522
18.5157 | -105.6576
18.5194 | -105.6648
18.5232 | -105.6712
18.5269 | -105.6744
18.5306 | -105.6775
18.5344 | -105.6817
18.5381 | -105.6866
18.5418 | -105.6896
18.5455 | -105.6903
18.5493 | -105.6903
18.553 | -105.6915
18.5567 | -105.6913
18.5604 | -105.6909
18.5642 | -105.69
18.5679 | -105.6892
18.5716 | -105.6879
18.5754 | -105.6855
18.5791 | -105.6833
18.5828 | -105.6804
18.5865 | -105.6752
18.5903 | -105.6701
18.594 | -105.6682
18.5977 | -105.6639
18.6014 | -105.6585
18.6052 | -105.6524
18.6089 | -105.6449
18.6126 | -105.6367
18.6164 | -105.6283
18.6201 | -105.6193
18.6238 | -105.6113
18.6275 | -105.6037
18.6313 | -105.596
18.635 | -105.5876
18.6387 | -105.5784
18.6424 | -105.5706
18.6462 | -105.5653
18.6499 | -105.5603
18.6536 | -105.5537
18.6573 | -105.5482
18.6611 | -105.5415
18.6648 | -105.5353
18.6685 | -105.5315
18.6723 | -105.5291
18.676 | -105.5295
18.6797 | -105.5276
18.6834 | -105.5268
18.6872 | -105.527
18.6909 | -105.5264
18.6946 | -105.5288
18.6983 | -105.5328
18.7021 | -105.5393
18.7058 | -105.5449
18.7095 | -105.5494
18.7133 | -105.5556
18.717 | -105.563
18.7207 | -105.571
18.7244 | -105.5799
18.7282 | -105.5901
18.7319 | -105.599
18.7356 | -105.6079
18.7393 | -105.6158
18.7431 | -105.6219
18.7468 | -105.6301
18.7505 | -105.6383
18.7543 | -105.6465
18.758 | -105.6533
18.7617 | -105.6587
18.7654 | -105.6635
18.7692 | -105.6673
18.7729 | -105.67
18.7766 | -105.676
18.7803 | -105.682
18.7841 | -105.685
18.7878 | -105.6879
18.7915 | -105.6922
18.7952 | -105.6945
18.799 | -105.6962
18.8027 | -105.6979
18.8064 | -105.6995
18.8102 | -105.7022
18.8139 | -105.7038
18.8176 | -105.6998
18.8213 | -105.7003
18.8251 | -105.7024
18.8288 | -105.7009
18.8325 | -105.6995
18.8362 | -105.6992
18.84 | -105.6978
18.8437 | -105.6949
18.8474 | -105.6923
18.8512 | -105.6889
18.8549 | -105.6871
18.8586 | -105.6851
18.8623 | -105.681
18.8661 | -105.6764
18.8698 | -105.6717
18.8735 | -105.6689
18.8772 | -105.6673
18.881 | -105.6658
18.8847 | -105.6639
18.8884 | -105.6617
18.8922 | -105.6595
18.8959 | -105.6596
18.8996 | -105.6582
18.9033 | -105.655
18.9071 | -105.6549
18.9108 | -105.6567
18.9145 | -105.6576
18.9182 | -105.6558
18.922 | -105.6549
18.9257 | -105.6542
18.9294 | -105.6544
18.9331 | -105.6571
18.9369 | -105.6593
18.9406 | -105.6607
18.9443 | -105.6626
18.9481 | -105.6653
18.9518 | -105.6705
18.9555 | -105.6755
18.9592 | -105.6808
18.963 | -105.6815
18.9667 | -105.6824
18.9704 | -105.6827
18.9741 | -105.6816
18.9779 | -105.6795
18.9816 | -105.6767
18.9853 | -105.675
18.9891 | -105.6737
18.9928 | -105.6714
18.9965 | -105.6676
19.0002 | -105.6615
19.004 | -105.656
19.0077 | -105.6527
19.0114 | -105.6493
19.0151 | -105.6413
19.0189 | -105.6353
19.0226 | -105.6274
19.0263 | -105.6203
19.03 | -105.6156
19.0338 | -105.6082
19.0375 | -105.6002
19.0412 | -105.5933
19.045 | -105.5838
19.0487 | -105.576
19.0524 | -105.5701
19.0561 | -105.5619
19.0599 | -105.5536
19.0636 | -105.5465
19.0673 | -105.5367
19.071 | -105.5258
19.0748 | -105.5173
19.0785 | -105.511
19.0822 | -105.5035
19.086 | -105.4957
19.0897 | -105.4901
19.0934 | -105.4841
19.0971 | -105.4781
19.1009 | -105.4735
19.1046 | -105.4694
19.1083 | -105.4657
19.112 | -105.461
19.1158 | -105.4574
19.1195 | -105.4555
19.1232 | -105.4548
19.127 | -105.4535
19.1307 | -105.4545
19.1344 | -105.4555
19.1381 | -105.4576
19.1419 | -105.4597
19.1456 | -105.4625
19.1493 | -105.4621
19.153 | -105.4621
19.1568 | -105.4645
19.1605 | -105.4672
19.1642 | -105.4705
19.1679 | -105.4729
19.1717 | -105.4757
19.1754 | -105.4807
19.1791 | -105.4848
19.1829 | -105.4879
19.1866 | -105.4903
19.1903 | -105.495
19.194 | -105.5021
19.1978 | -105.5066
19.2015 | -105.5108
19.2052 | -105.511
19.2089 | -105.5111
19.2127 | -105.5131
19.2164 | -105.5153
19.2201 | -105.5182
19.2239 | -105.5162
19.2276 | -105.5159
19.2313 | -105.5158
19.235 | -105.5168
19.2388 | -105.5177
19.2425 | -105.5188
19.2462 | -105.5229
19.2499 | -105.5257
19.2537 | -105.527
19.2574 | -105.5269
19.2611 | -105.5238
19.2648 | -105.5213
19.2686 | -105.5218
19.2723 | -105.5245
19.276 | -105.5226
19.2798 | -105.518
19.2835 | -105.5149
19.2872 | -105.5147
19.2909 | -105.5141
19.2947 | -105.5159
19.2984 | -105.5147
19.3021 | -105.5124
19.3058 | -105.514
19.3096 | -105.5164
19.3133 | -105.5179
19.317 | -105.5184
19.3208 | -105.5176
19.3245 | -105.5169
19.3282 | -105.5177
19.3319 | -105.5179
19.3357 | -105.517
19.3394 | -105.5183
19.3431 | -105.5212
19.3468 | -105.5217
19.3506 | -105.5206
19.3543 | -105.5203
19.358 | -105.5199
19.3618 | -105.5217
19.3655 | -105.5211
19.3692 | -105.5192
19.3729 | -105.5159
19.3767 | -105.5126
19.3804 | -105.5101
19.3841 | -105.5076
19.3878 | -105.5061
19.3916 | -105.5032
19.3953 | -105.4984
19.399 | -105.4949
19.4027 | -105.4925
19.4065 | -105.488
19.4102 | -105.4831
19.4139 | -105.4786
19.4177 | -105.4764
19.4214 | -105.4722
19.4251 | -105.4697
19.4288 | -105.4673
19.4326 | -105.463
19.4363 | -105.4584
19.44 | -105.4536
19.4437 | -105.4493
19.4475 | -105.443
19.4512 | -105.4364
19.4549 | -105.4311
19.4587 | -105.4256
19.4624 | -105.4198
19.4661 | -105.4154
19.4698 | -105.4085
19.4736 | -105.4037
19.4773 | -105.3984
19.481 | -105.3931
19.4847 | -105.3889
19.4885 | -105.3834
19.4922 | -105.3764
19.4959 | -105.3717
19.4997 | -105.3701
19.5034 | -105.3652
19.5071 | -105.3621
19.5108 | -105.3618
19.5146 | -105.3594
19.5183 | -105.3559
19.522 | -105.3519
19.5257 | -105.3498
19.5295 | -105.35
19.5332 | -105.3521
19.5369 | -105.3544
19.5406 | -105.3561
19.5444 | -105.3574
19.5481 | -105.3593
19.5518 | -105.3655
19.5556 | -105.3725
19.5593 | -105.3768
19.563 | -105.3821
19.5667 | -105.3858
19.5705 | -105.386
19.5742 | -105.3908
19.5779 | -105.3964
19.5816 | -105.4026
19.5854 | -105.4066
19.5891 | -105.4125
19.5928 | -105.4187
19.5966 | -105.4268
19.6003 | -105.4346
19.604 | -105.4417
19.6077 | -105.4486
19.6115 | -105.4569
19.6152 | -105.4638
19.6189 | -105.4693
19.6226 | -105.4761
19.6264 | -105.4814
19.6301 | -105.4859
19.6338 | -105.4912
19.6375 | -105.496
19.6413 | -105.4994
19.645 | -105.5023
19.6487 | -105.5073
19.6525 | -105.5136
19.6562 | -105.5198
19.6599 | -105.5229
19.6636 | -105.5234
19.6674 | -105.5237
19.6711 | -105.5247
19.6748 | -105.5259
19.6785 | -105.5252
19.6823 | -105.5244
19.686 | -105.5228
19.6897 | -105.521
19.6935 | -105.5186
19.6972 | -105.5143
19.7009 | -105.5134
19.7046 | -105.5135
19.7084 | -105.5116
19.7121 | -105.5101
19.7158 | -105.5086
19.7195 | -105.5064
19.7233 | -105.5039
19.727 | -105.5045
19.7307 | -105.5045
19.7345 | -105.5018
19.7382 | -105.5026
19.7419 | -105.5013
19.7456 | -105.4982
19.7494 | -105.4973
19.7531 | -105.497
19.7568 | -105.4953
19.7605 | -105.4926
19.7643 | -105.4873
19.768 | -105.4826
19.7717 | -105.4796
19.7754 | -105.4753
19.7792 | -105.4713
19.7829 | -105.4693
19.7866 | -105.4646
19.7904 | -105.4567
19.7941 | -105.4502
19.7978 | -105.4441
19.8015 | -105.4391
19.8053 | -105.433
19.809 | -105.4234
19.8127 | -105.4118
19.8164 | -105.4011
19.8202 | -105.3914
19.8239 | -105.3799
19.8276 | -105.3698
19.8314 | -105.3593
19.8351 | -105.348
19.8388 | -105.3376
19.8425 | -105.3271
19.8463 | -105.319
19.85 | -105.3118
19.8537 | -105.3031
19.8574 | -105.2946
19.8612 | -105.2845
19.8649 | -105.2777
19.8686 | -105.2699
19.8724 | -105.2629
19.8761 | -105.2545
19.8798 | -105.247
19.8835 | -105.2427
19.8873 | -105.2395
19.891 | -105.2426
19.8947 | -105.2405
19.8984 | -105.2408
19.9022 | -105.2424
19.9059 | -105.2435
19.9096 | -105.2454
19.9133 | -105.2486
19.9171 | -105.255
19.9208 | -105.2614
19.9245 | -105.2687
19.9283 | -105.2794
19.932 | -105.2881
19.9357 | -105.2976
19.9394 | -105.3105
19.9432 | -105.326
19.9469 | -105.3413
19.9506 | -105.3553
19.9543 | -105.371
19.9581 | -105.3876
19.9618 | -105.4035
19.9655 | -105.4204
19.9693 | -105.4356
19.973 | -105.4506
19.9767 | -105.4669
19.9804 | -105.4827
19.9842 | -105.4968
19.9879 | -105.5101
19.9916 | -105.5231
19.9953 | -105.5345
19.9991 | -105.5442
20.0028 | -105.5498
20.0065 | -105.5541
20.0102 | -105.5569
20.014 | -105.5579
20.0177 | -105.5582
20.0214 | -105.5529
20.0252 | -105.5487
20.0289 | -105.5446
20.0326 | -105.5366
20.0363 | -105.5262
20.0401 | -105.5151
20.0438 | -105.5035
20.0475 | -105.4898
20.0512 | -105.4741
20.055 | -105.4557
20.0587 | -105.4349
20.0624 | -105.4111
20.0662 | -105.3868
20.0699 | -105.3634
20.0736 | -105.3403
20.0773 | -105.3159
20.0811 | -105.2934
20.0848 | -105.2684
20.0885 | -105.2457
20.0922 | -105.2231
20.096 | -105.2003
20.0997 | -105.178
20.1034 | -105.156
20.1072 | -105.1364
20.1109 | -105.1148
20.1146 | -105.0942
20.1183 | -105.0753
20.1221 | -105.0568
20.1258 | -105.0405
20.1295 | -105.0248
20.1332 | -105.0122
20.137 | -104.9977
20.1407 | -104.9846
20.1444 | -104.975
20.1481 | -104.9673
20.1519 | -104.9599
20.1556 | -104.9534
20.1593 | -104.9501
20.1631 | -104.9474
20.1668 | -104.9443
20.1705 | -104.9416
20.1742 | -104.9417
20.178 | -104.944
20.1817 | -104.9457
20.1854 | -104.9459
20.1891 | -104.9476
20.1929 | -104.9524
20.1966 | -104.9569
20.2003 | -104.962
20.2041 | -104.9656
20.2078 | -104.9699
20.2115 | -104.9725
20.2152 | -104.9743
20.219 | -104.9766
20.2227 | -104.9794
20.2264 | -104.983
20.2301 | -104.987
20.2339 | -104.9937
20.2376 | -105.0012
20.2413 | -105.0062
20.245 | -105.0092
20.2488 | -105.0143
20.2525 | -105.0189
20.2562 | -105.0256
20.26 | -105.0297
20.2637 | -105.0333
20.2674 | -105.0379
20.2711 | -105.0407
20.2749 | -105.0431
20.2786 | -105.0463
20.2823 | -105.0484
20.286 | -105.0511
20.2898 | -105.055
20.2935 | -105.058
20.2972 | -105.0624
20.301 | -105.0665
20.3047 | -105.0696
20.3084 | -105.0727
20.3121 | -105.0762
20.3159 | -105.078
20.3196 | -105.081
20.3233 | -105.0829
20.327 | -105.0835
20.3308 | -105.0863
20.3345 | -105.0881
20.3382 | -105.0897
20.342 | -105.0926
20.3457 | -105.0975
20.3494 | -105.1028
20.3531 | -105.1052
20.3569 | -105.1054
20.3606 | -105.1092
20.3643 | -105.1137
20.368 | -105.1168
20.3718 | -105.1207
20.3755 | -105.1227
20.3792 | -105.1273
20.3829 | -105.1337
20.3867 | -105.1383
20.3904 | -105.1416
20.3941 | -105.1452
20.3979 | -105.1495
20.4016 | -105.1533
20.4053 | -105.1567
20.409 | -105.1591
20.4128 | -105.1612
20.4165 | -105.164
20.4202 | -105.1657
20.4239 | -105.1692
20.4277 | -105.1723
20.4314 | -105.1713
20.4351 | -105.1724
20.4389 | -105.1728
20.4426 | -105.1739
20.4463 | -105.1755
20.45 | -105.1753
20.4538 | -105.1727
20.4575 | -105.17
20.4612 | -105.1693
20.4649 | -105.1668
20.4687 | -105.1634
20.4724 | -105.1601
20.4761 | -105.1583
20.4799 | -105.1554
20.4836 | -105.1529
20.4873 | -105.1501
20.491 | -105.1478
20.4948 | -105.1459
20.4985 | -105.142
20.5022 | -105.139
20.5059 | -105.136
20.5097 | -105.1341
20.5134 | -105.1316
20.5171 | -105.1269
20.5208 | -105.1223
20.5246 | -105.1174
20.5283 | -105.1137
20.532 | -105.1101
20.5358 | -105.1067
20.5395 | -105.1
20.5432 | -105.0938
20.5469 | -105.0902
20.5507 | -105.0881
20.5544 | -105.0853
20.5581 | -105.0792
20.5618 | -105.0735
20.5656 | -105.0701
20.5693 | -105.0671
20.573 | -105.0654
20.5768 | -105.0617
20.5805 | -105.0574
20.5842 | -105.0541
20.5879 | -105.0486
20.5917 | -105.0427
20.5954 | -105.038
20.5991 | -105.0348
20.6028 | -105.032
20.6066 | -105.0276
20.6103 | -105.022
20.614 | -105.0181
20.6178 | -105.0156
20.6215 | -105.011
20.6252 | -105.0082
20.6289 | -105.0052
20.6327 | -105.0002
20.6364 | -104.9938
20.6401 | -104.9898
20.6438 | -104.9875
20.6476 | -104.9859
20.6513 | -104.9792
20.655 | -104.9713
20.6587 | -104.9618
20.6625 | -104.9535
20.6662 | -104.9481
20.6699 | -104.9408
20.6737 | -104.9323
20.6774 | -104.9244
20.6811 | -104.9158
20.6848 | -104.9049
20.6886 | -104.8973
20.6923 | -104.8895
20.696 | -104.8803
20.6997 | -104.873
20.7035 | -104.8625
20.7072 | -104.8497
20.7109 | -104.8372
20.7147 | -104.824
20.7184 | -104.8109
20.7221 | -104.8016
20.7258 | -104.7904
20.7296 | -104.7796
20.7333 | -104.7682
20.737 | -104.7566
20.7407 | -104.7447
20.7445 | -104.7348
20.7482 | -104.7271
20.7519 | -104.7174
20.7556 | -104.7061
20.7594 | -104.697
20.7631 | -104.6887
20.7668 | -104.6811
20.7706 | -104.6728
20.7743 | -104.6635
20.778 | -104.6519
20.7817 | -104.6423
20.7855 | -104.6343
20.7892 | -104.6291
20.7929 | -104.6245
20.7966 | -104.6183
20.8004 | -104.6112
20.8041 | -104.6044
20.8078 | -104.5994
20.8116 | -104.5948
20.8153 | -104.5907
20.819 | -104.5849
20.8227 | -104.5807
20.8265 | -104.5787
20.8302 | -104.5761
20.8339 | -104.5741
20.8376 | -104.5733
20.8414 | -104.5711
20.8451 | -104.5687
20.8488 | -104.5658
20.8526 | -104.5655
20.8563 | -104.5668
20.86 | -104.5678
20.8637 | -104.5689
20.8675 | -104.5705
20.8712 | -104.5733
20.8749 | -104.5775
20.8786 | -104.5801
20.8824 | -104.5829
20.8861 | -104.589
20.8898 | -104.595
20.8935 | -104.5998
20.8973 | -104.605
20.901 | -104.6124
20.9047 | -104.62
20.9085 | -104.6282
20.9122 | -104.6338
20.9159 | -104.6412
20.9196 | -104.648
20.9234 | -104.6537
20.9271 | -104.6608
20.9308 | -104.6674
20.9345 | -104.6719
20.9383 | -104.6776
20.942 | -104.6841
20.9457 | -104.6912
20.9495 | -104.697
20.9532 | -104.7046
20.9569 | -104.7096
20.9606 | -104.7125
20.9644 | -104.7162
20.9681 | -104.7139
20.9718 | -104.7152
20.9755 | -104.7174
20.9793 | -104.7192
20.983 | -104.7221
20.9867 | -104.7219
20.9904 | -104.7217
20.9942 | -104.7221
20.9979 | -104.7211
21.0016 | -104.7212
21.0054 | -104.7209
21.0091 | -104.7197
21.0128 | -104.7188
21.0165 | -104.7175
21.0203 | -104.7134
21.024 | -104.7115
21.0277 | -104.7083
21.0314 | -104.705
21.0352 | -104.7017
21.0389 | -104.6978
21.0426 | -104.6963
21.0464 | -104.6941
21.0501 | -104.6918
21.0538 | -104.6905
21.0575 | -104.6883
21.0613 | -104.6862
21.065 | -104.6855
21.0687 | -104.6848
21.0724 | -104.6876
21.0762 | -104.6919
21.0799 | -104.6949
21.0836 | -104.6951
21.0874 | -104.6949
21.0911 | -104.6986
21.0948 | -104.7027
21.0985 | -104.706
21.1023 | -104.7095
21.106 | -104.7114
21.1097 | -104.7134
21.1134 | -104.7169
21.1172 | -104.7241
21.1209 | -104.731
21.1246 | -104.7356
21.1283 | -104.7399
21.1321 | -104.7476
21.1358 | -104.7543
21.1395 | -104.7606
21.1433 | -104.769
21.147 | -104.7761
21.1507 | -104.7808
21.1544 | -104.7835
21.1582 | -104.7843
21.1619 | -104.7847
21.1656 | -104.7869
21.1693 | -104.7916
21.1731 | -104.7956
21.1768 | -104.7994
21.1805 | -104.8013
21.1843 | -104.8015
21.188 | -104.8027
21.1917 | -104.8049
21.1954 | -104.8062
21.1992 | -104.8068
21.2029 | -104.8069
21.2066 | -104.8035
21.2103 | -104.8036
21.2141 | -104.803
21.2178 | -104.8023
21.2215 | -104.8015
21.2252 | -104.7998
21.229 | -104.7974
21.2327 | -104.795
21.2364 | -104.7919
21.2402 | -104.7878
21.2439 | -104.7832
21.2476 | -104.7805
21.2513 | -104.7801
21.2551 | -104.7796
21.2588 | -104.7774
21.2625 | -104.7741
21.2662 | -104.7746
21.27 | -104.7764
21.2737 | -104.7774
21.2774 | -104.7782
21.2812 | -104.7805
21.2849 | -104.7815
21.2886 | -104.7822
21.2923 | -104.7847
21.2961 | -104.7885
21.2998 | -104.7945
21.3035 | -104.801
21.3072 | -104.8081
21.311 | -104.8149
21.3147 | -104.821
21.3184 | -104.829
21.3222 | -104.8383
21.3259 | -104.8474
21.3296 | -104.859
21.3333 | -104.8712
21.3371 | -104.8837
21.3408 | -104.8977
21.3445 | -104.9094
21.3482 | -104.9222
21.352 | -104.9355
21.3557 | -104.9476
21.3594 | -104.9585
21.3631 | -104.9701
21.3669 | -104.9805
21.3706 | -104.9924
21.3743 | -105.0026
21.3781 | -105.0127
21.3818 | -105.0223
21.3855 | -105.0349
21.3892 | -105.0438
21.393 | -105.0507
21.3967 | -105.0608
21.4004 | -105.0707
21.4041 | -105.079
21.4079 | -105.0869
21.4116 | -105.0941
21.4153 | -105.0998
21.4191 | -105.104
21.4228 | -105.1073
21.4265 | -105.1106
21.4302 | -105.1127
21.434 | -105.1109
21.4377 | -105.1087
21.4414 | -105.1081
21.4451 | -105.1057
21.4489 | -105.1053
21.4526 | -105.105
21.4563 | -105.1055
21.4601 | -105.1045
21.4638 | -105.1029
21.4675 | -105.0984
21.4712 | -105.0916
21.475 | -105.0863
21.4787 | -105.0807
21.4824 | -105.0766
21.4861 | -105.0701
21.4899 | -105.0624
21.4936 | -105.0547
21.4973 | -105.0455
21.501 | -105.0379
21.5048 | -105.0323
21.5085 | -105.025
21.5122 | -105.0208
21.516 | -105.0158
21.5197 | -105.0083
21.5234 | -105.0005
21.5271 | -104.9954
21.5309 | -104.9915
21.5346 | -104.9884
21.5383 | -104.9838
21.542 | -104.9772
21.5458 | -104.972
21.5495 | -104.9692
21.5532 | -104.965
21.557 | -104.9576
21.5607 | -104.951
21.5644 | -104.9474
21.5681 | -104.9445
21.5719 | -104.9398
21.5756 | -104.9367
21.5793 | -104.9326
21.583 | -104.927
21.5868 | -104.9234
21.5905 | -104.9214
21.5942 | -104.9195
21.598 | -104.9149
21.6017 | -104.9107
21.6054 | -104.907
21.6091 | -104.9025
21.6129 | -104.9
21.6166 | -104.9002
21.6203 | -104.9008
21.624 | -104.9011
21.6278 | -104.8991
21.6315 | -104.8963
21.6352 | -104.8949
21.6389 | -104.8928
21.6427 | -104.8916
21.6464 | -104.8906
21.6501 | -104.8888
21.6539 | -104.8867
21.6576 | -104.8876
21.6613 | -104.8886
21.665 | -104.8883
21.6688 | -104.889
21.6725 | -104.8918
21.6762 | -104.8927
21.6799 | -104.8942
21.6837 | -104.8964
21.6874 | -104.8975
21.6911 | -104.897
21.6949 | -104.9012
21.6986 | -104.9043
21.7023 | -104.9059
21.706 | -104.9069
21.7098 | -104.907
21.7135 | -104.9075
21.7172 | -104.9055
21.7209 | -104.9046
21.7247 | -104.9043
21.7284 | -104.9028
21.7321 | -104.903
21.7358 | -104.9058
21.7396 | -104.9082
21.7433 | -104.9086
21.747 | -104.9062
21.7508 | -104.904
21.7545 | -104.9018
21.7582 | -104.8995
21.7619 | -104.8961
21.7657 | -104.8931
21.7694 | -104.8905
21.7731 | -104.8884
21.7768 | -104.8847
21.7806 | -104.8788
21.7843 | -104.8748
21.788 | -104.8712
21.7918 | -104.8697
21.7955 | -104.8676
21.7992 | -104.8649
21.8029 | -104.8641
21.8067 | -104.8609
21.8104 | -104.855
21.8141 | -104.8513
21.8178 | -104.8492
21.8216 | -104.8456
21.8253 | -104.8412
21.829 | -104.8384
21.8328 | -104.8359
21.8365 | -104.8317
21.8402 | -104.8259
21.8439 | -104.8225
21.8477 | -104.8182
21.8514 | -104.8155
21.8551 | -104.8124
21.8588 | -104.8066
21.8626 | -104.8009
21.8663 | -104.7963
21.87 | -104.7917
21.8737 | -104.787
21.8775 | -104.7814
21.8812 | -104.7744
21.8849 | -104.7674
21.8887 | -104.7604
21.8924 | -104.7532
21.8961 | -104.7455
21.8998 | -104.737
21.9036 | -104.7284
21.9073 | -104.7217
21.911 | -104.7148
21.9147 | -104.7064
21.9185 | -104.698
21.9222 | -104.6876
21.9259 | -104.6767
21.9297 | -104.6674
21.9334 | -104.6579
21.9371 | -104.6493
21.9408 | -104.6418
21.9446 | -104.6352
21.9483 | -104.6307
21.952 | -104.6243
21.9557 | -104.6165
21.9595 | -104.609
21.9632 | -104.6038
21.9669 | -104.5972
21.9706 | -104.591
21.9744 | -104.5859
21.9781 | -104.5801
21.9818 | -104.5744
21.9856 | -104.57
21.9893 | -104.5638
21.993 | -104.5589
21.9967 | -104.5574
22.0005 | -104.5526
22.0042 | -104.5467
22.0079 | -104.5424
22.0116 | -104.5385
22.0154 | -104.5347
22.0191 | -104.532
22.0228 | -104.5286
22.0266 | -104.5282
22.0303 | -104.5256
22.034 | -104.5214
22.0377 | -104.5193
22.0415 | -104.5167
22.0452 | -104.5108
22.0489 | -104.5043
22.0526 | -104.4981
22.0564 | -104.4912
22.0601 | -104.4857
22.0638 | -104.4808
22.0676 | -104.4749
22.0713 | -104.4697
22.075 | -104.4635
22.0787 | -104.4573
22.0825 | -104.4505
22.0862 | -104.4446
22.0899 | -104.4396
22.0936 | -104.4325
22.0974 | -104.4249
22.1011 | -104.4174
22.1048 | -104.4085
22.1085 | -104.4001
22.1123 | -104.3938
22.116 | -104.3875
22.1197 | -104.3826
22.1235 | -104.3779
22.1272 | -104.3726
22.1309 | -104.3671
22.1346 | -104.3615
22.1384 | -104.3576
22.1421 | -104.3541
22.1458 | -104.3507
22.1495 | -104.3472
22.1533 | -104.3452
22.157 | -104.3423
22.1607 | -104.3425
22.1645 | -104.3426
22.1682 | -104.3423
22.1719 | -104.3437
22.1756 | -104.3457
22.1794 | -104.3483
22.1831 | -104.3523
22.1868 | -104.3549
22.1905 | -104.3595
22.1943 | -104.3658
22.198 | -104.3741
22.2017 | -104.3841
22.2055 | -104.3944
22.2092 | -104.4036
22.2129 | -104.412
22.2166 | -104.4231
22.2204 | -104.4365
22.2241 | -104.449
22.2278 | -104.4608
22.2315 | -104.4738
22.2353 | -104.4883
22.239 | -104.5011
22.2427 | -104.5129
22.2464 | -104.5257
22.2502 | -104.5386
22.2539 | -104.5525
22.2576 | -104.5635
22.2614 | -104.5734
22.2651 | -104.5863
22.2688 | -104.6016
22.2725 | -104.6164
22.2763 | -104.6272
22.28 | -104.6387
22.2837 | -104.6492
22.2874 | -104.6588
22.2912 | -104.6686
22.2949 | -104.6766
22.2986 | -104.6836
22.3024 | -104.6912
22.3061 | -104.6963
22.3098 | -104.7
22.3135 | -104.7066
22.3173 | -104.7115
22.321 | -104.7143
22.3247 | -104.7173
22.3284 | -104.7212
22.3322 | -104.723
22.3359 | -104.7238
22.3396 | -104.7238
22.3434 | -104.7242
22.3471 | -104.7242
22.3508 | -104.7229
22.3545 | -104.7217
22.3583 | -104.7224
22.362 | -104.7239
22.3657 | -104.725
22.3694 | -104.7259
22.3732 | -104.727
22.3769 | -104.7269
22.3806 | -104.7266
22.3843 | -104.7263
22.3881 | -104.7258
22.3918 | -104.7282
22.3955 | -104.7309
22.3993 | -104.7311
22.403 | -104.7299
22.4067 | -104.7302
22.4104 | -104.7345
22.4142 | -104.7382
22.4179 | -104.7413
22.4216 | -104.7476
22.4253 | -104.7526
22.4291 | -104.7546
22.4328 | -104.7567
22.4365 | -104.7588
22.4403 | -104.7634
22.444 | -104.7675
22.4477 | -104.7697
22.4514 | -104.7728
22.4552 | -104.7747
22.4589 | -104.7769
22.4626 | -104.7798
22.4663 | -104.7812
22.4701 | -104.7802
22.4738 | -104.7806
22.4775 | -104.7803
22.4812 | -104.778
22.485 | -104.7752
22.4887 | -104.7728
22.4924 | -104.7694
22.4962 | -104.7626
22.4999 | -104.7538
22.5036 | -104.7455
22.5073 | -104.7357
22.5111 | -104.7224
22.5148 | -104.7044
22.5185 | -104.6866
22.5222 | -104.6687
22.526 | -104.649
22.5297 | -104.6288
22.5334 | -104.6055
22.5372 | -104.5809
22.5409 | -104.5573
22.5446 | -104.5343
22.5483 | -104.5106
22.5521 | -104.4842
22.5558 | -104.4575
22.5595 | -104.43
22.5632 | -104.4003
22.567 | -104.371
22.5707 | -104.3398
22.5744 | -104.3101
22.5782 | -104.2798
22.5819 | -104.2421
22.5856 | -104.2093
22.5893 | -104.1759
22.5931 | -104.1448
22.5968 | -104.1163
22.6005 | -104.0874
22.6042 | -104.0596
22.608 | -104.0325
22.6117 | -104.0048
22.6154 | -103.9777
22.6191 | -103.9542
22.6229 | -103.9322
22.6266 | -103.9117
22.6303 | -103.8926
22.6341 | -103.8756
22.6378 | -103.8611
22.6415 | -103.8484
22.6452 | -103.8362
22.649 | -103.8247
22.6527 | -103.8156
22.6564 | -103.8112
22.6601 | -103.808
22.6639 | -103.8069
22.6676 | -103.8084
22.6713 | -103.812
22.6751 | -103.8171
22.6788 | -103.8236
22.6825 | -103.833
22.6862 | -103.8455
22.69 | -103.8593
22.6937 | -103.8757
22.6974 | -103.8945
22.7011 | -103.914
22.7049 | -103.9365
22.7086 | -103.9606
22.7123 | -103.9858
22.716 | -104.0125
22.7198 | -104.0411
22.7235 | -104.0714
22.7272 | -104.1029
22.731 | -104.1338
22.7347 | -104.1658
22.7384 | -104.1986
22.7421 | -104.2306
22.7459 | -104.2646
22.7496 | -104.3007
22.7533 | -104.3367
22.757 | -104.3713
22.7608 | -104.4056
22.7645 | -104.4399
22.7682 | -104.4732
22.772 | -104.5045
22.7757 | -104.5356
22.7794 | -104.5659
22.7831 | -104.5942
22.7869 | -104.6236
22.7906 | -104.6517
22.7943 | -104.6754
22.798 | -104.698
22.8018 | -104.7212
22.8055 | -104.7425
22.8092 | -104.7603
22.813 | -104.7774
22.8167 | -104.794
22.8204 | -104.8102
22.8241 | -104.8256
22.8279 | -104.8383
22.8316 | -104.8504
22.8353 | -104.8633
22.839 | -104.8734
22.8428 | -104.8804
22.8465 | -104.8869
22.8502 | -104.8936
22.8539 | -104.8988
22.8577 | -104.9012
22.8614 | -104.9017
22.8651 | -104.9002
22.8689 | -104.8964
22.8726 | -104.8953
22.8763 | -104.8948
22.88 | -104.8937
22.8838 | -104.8924
22.8875 | -104.8896
22.8912 | -104.8875
22.8949 | -104.8826
22.8987 | -104.8791
22.9024 | -104.8794
22.9061 | -104.88
22.9099 | -104.8777
22.9136 | -104.8738
22.9173 | -104.8698
22.921 | -104.8675
22.9248 | -104.8656
22.9285 | -104.8635
22.9322 | -104.8607
22.9359 | -104.858
22.9397 | -104.8572
22.9434 | -104.8556
22.9471 | -104.8553
22.9508 | -104.8559
22.9546 | -104.8561
22.9583 | -104.8564
22.962 | -104.8565
22.9658 | -104.8586
22.9695 | -104.8625
22.9732 | -104.8676
22.9769 | -104.8732
22.9807 | -104.8756
22.9844 | -104.8779
22.9881 | -104.8806
22.9918 | -104.8821
22.9956 | -104.8846
22.9993 | -104.888
23.003 | -104.8933
23.0068 | -104.8977
23.0105 | -104.9006
23.0142 | -104.9046
23.0179 | -104.9093
23.0217 | -104.9132
23.0254 | -104.9172
23.0291 | -104.9184
23.0328 | -104.9197
23.0366 | -104.9196
23.0403 | -104.9192
23.044 | -104.92
23.0478 | -104.921
23.0515 | -104.9207
23.0552 | -104.9203
23.0589 | -104.9193
23.0627 | -104.9181
23.0664 | -104.9165
23.0701 | -104.9158
23.0738 | -104.9149
23.0776 | -104.9132
23.0813 | -104.9126
23.085 | -104.9105
23.0887 | -104.9084
23.0925 | -104.9086
23.0962 | -104.9098
23.0999 | -104.9092
23.1037 | -104.9074
23.1074 | -104.9066
23.1111 | -104.9061
23.1148 | -104.9076
23.1186 | -104.9115
23.1223 | -104.9137
23.126 | -104.9149
23.1297 | -104.9156
23.1335 | -104.9177
23.1372 | -104.9207
23.1409 | -104.9231
23.1447 | -104.9275
23.1484 | -104.9312
23.1521 | -104.9331
23.1558 | -104.937
23.1596 | -104.942
23.1633 | -104.946
23.167 | -104.9503
23.1707 | -104.9547
23.1745 | -104.958
23.1782 | -104.96
23.1819 | -104.9634
23.1857 | -104.9656
23.1894 | -104.966
23.1931 | -104.9699
23.1968 | -104.9746
23.2006 | -104.9761
23.2043 | -104.9767
23.208 | -104.9769
23.2117 | -104.9789
23.2155 | -104.9797
23.2192 | -104.9791
23.2229 | -104.9799
23.2266 | -104.9784
23.2304 | -104.9744
23.2341 | -104.97
23.2378 | -104.9678
23.2416 | -104.9651
23.2453 | -104.9611
23.249 | -104.955
23.2527 | -104.9498
23.2565 | -104.9454
23.2602 | -104.9419
23.2639 | -104.9372
23.2676 | -104.931
23.2714 | -104.9246
23.2751 | -104.9185
23.2788 | -104.9109
23.2826 | -104.904
23.2863 | -104.8968
23.29 | -104.8898
23.2937 | -104.884
23.2975 | -104.8779
23.3012 | -104.8706
23.3049 | -104.8649
23.3086 | -104.8597
23.3124 | -104.856
23.3161 | -104.8536
23.3198 | -104.8518
23.3236 | -104.8485
23.3273 | -104.8459
23.331 | -104.8429
23.3347 | -104.8416
23.3385 | -104.8393
23.3422 | -104.8386
23.3459 | -104.838
23.3496 | -104.8388
23.3534 | -104.8413
23.3571 | -104.8464
23.3608 | -104.8515
23.3645 | -104.8559
23.3683 | -104.8598
23.372 | -104.864
23.3757 | -104.8681
23.3795 | -104.8742
23.3832 | -104.8825
23.3869 | -104.8911
23.3906 | -104.8981
23.3944 | -104.9057
23.3981 | -104.912
23.4018 | -104.9196
23.4055 | -104.9296
23.4093 | -104.9406
23.413 | -104.9497
23.4167 | -104.9589
23.4205 | -104.9672
23.4242 | -104.9746
23.4279 | -104.9824
23.4316 | -104.9909
23.4354 | -104.9989
23.4391 | -105.0049
23.4428 | -105.0106
23.4465 | -105.0163
23.4503 | -105.0208
23.454 | -105.0247
23.4577 | -105.0277
23.4614 | -105.031
23.4652 | -105.0337
23.4689 | -105.0345
23.4726 | -105.0343
23.4764 | -105.033
23.4801 | -105.0314
23.4838 | -105.0272
23.4875 | -105.0221
23.4913 | -105.019
23.495 | -105.016
23.4987 | -105.0119
23.5024 | -105.0085
23.5062 | -105.0037
23.5099 | -104.9981
23.5136 | -104.992
23.5174 | -104.9863
23.5211 | -104.9804
23.5248 | -104.9722
23.5285 | -104.963
23.5323 | -104.955
23.536 | -104.947
23.5397 | -104.9381
23.5434 | -104.928
23.5472 | -104.9196
23.5509 | -104.9131
23.5546 | -104.9069
23.5584 | -104.902
23.5621 | -104.8964
23.5658 | -104.8924
23.5695 | -104.8885
23.5733 | -104.8851
23.577 | -104.8831
23.5807 | -104.8799
23.5844 | -104.8756
23.5882 | -104.8735
23.5919 | -104.8718
23.5956 | -104.8724
23.5993 | -104.8744
23.6031 | -104.8772
23.6068 | -104.8794
23.6105 | -104.8836
23.6143 | -104.8897
23.618 | -104.8967
23.6217 | -104.9051
23.6254 | -104.9126
23.6292 | -104.9212
23.6329 | -104.9314
23.6366 | -104.9404
23.6403 | -104.9491
23.6441 | -104.9576
23.6478 | -104.966
23.6515 | -104.9776
23.6553 | -104.9919
23.659 | -105.0034
23.6627 | -105.0131
23.6664 | -105.0221
23.6702 | -105.0307
23.6739 | -105.0418
23.6776 | -105.0512
23.6813 | -105.0596
23.6851 | -105.0682
23.6888 | -105.0758
23.6925 | -105.0815
23.6962 | -105.0879
23.7 | -105.0956
23.7037 | -105.1029
23.7074 | -105.1079
23.7112 | -105.1109
23.7149 | -105.1105
23.7186 | -105.1099
23.7223 | -105.1082
23.7261 | -105.1067
23.7298 | -105.1046
23.7335 | -105.1019
23.7372 | -105.0958
23.741 | -105.0879
23.7447 | -105.0797
23.7484 | -105.0723
23.7522 | -105.0631
23.7559 | -105.0536
23.7596 | -105.0444
23.7633 | -105.0348
23.7671 | -105.0227
23.7708 | -105.0087
23.7745 | -104.995
23.7782 | -104.9816
23.782 | -104.9686
23.7857 | -104.9576
23.7894 | -104.9443
23.7932 | -104.931
23.7969 | -104.917
23.8006 | -104.9016
23.8043 | -104.8888
23.8081 | -104.8769
23.8118 | -104.8644
23.8155 | -104.8522
23.8192 | -104.8407
23.823 | -104.8295
23.8267 | -104.8172
23.8304 | -104.8056
23.8341 | -104.7945
23.8379 | -104.7842
23.8416 | -104.7771
23.8453 | -104.7706
23.8491 | -104.7635
23.8528 | -104.7548
23.8565 | -104.7478
23.8602 | -104.7421
23.864 | -104.7378
23.8677 | -104.7336
23.8714 | -104.7302
23.8751 | -104.7282
23.8789 | -104.727
23.8826 | -104.7248
23.8863 | -104.722
23.8901 | -104.7205
23.8938 | -104.7188
23.8975 | -104.7177
23.9012 | -104.7177
23.905 | -104.7181
23.9087 | -104.7194
23.9124 | -104.721
23.9161 | -104.7244
23.9199 | -104.7295
23.9236 | -104.734
23.9273 | -104.7392
23.9311 | -104.7446
23.9348 | -104.7496
23.9385 | -104.7546
23.9422 | -104.7598
23.946 | -104.7652
23.9497 | -104.7689
23.9534 | -104.772
23.9571 | -104.7757
23.9609 | -104.7806
23.9646 | -104.7861
23.9683 | -104.7911
23.972 | -104.7973
23.9758 | -104.8042
23.9795 | -104.8113
23.9832 | -104.8171
23.987 | -104.8218
23.9907 | -104.825
23.9944 | -104.8284
23.9981 | -104.8327
24.0019 | -104.8379
24.0056 | -104.8419
24.0093 | -104.8473
24.013 | -104.8523
24.0168 | -104.8553
24.0205 | -104.8566
24.0242 | -104.8584
24.028 | -104.86
24.0317 | -104.8604
24.0354 | -104.8612
24.0391 | -104.8625
24.0429 | -104.8648
24.0466 | -104.8677
24.0503 | -104.8705
24.054 | -104.8738
24.0578 | -104.8747
24.0615 | -104.8752
24.0652 | -104.8751
24.0689 | -104.8749
24.0727 | -104.8747
24.0764 | -104.8756
24.0801 | -104.8764
24.0839 | -104.8763
24.0876 | -104.8777
24.0913 | -104.8798
24.095 | -104.8802
24.0988 | -104.8802
24.1025 | -104.8792
24.1062 | -104.8789
24.1099 | -104.8784
24.1137 | -104.8774
24.1174 | -104.8753
24.1211 | -104.8717
24.1249 | -104.8677
24.1286 | -104.8646
24.1323 | -104.8626
24.136 | -104.8602
24.1398 | -104.8596
24.1435 | -104.8576
24.1472 | -104.8528
24.1509 | -104.8485
24.1547 | -104.8455
24.1584 | -104.8449
24.1621 | -104.8423
24.1659 | -104.838
24.1696 | -104.8333
24.1733 | -104.8291
24.177 | -104.8236
24.1808 | -104.8157
24.1845 | -104.8084
24.1882 | -104.8018
24.1919 | -104.7972
24.1957 | -104.793
24.1994 | -104.7888
24.2031 | -104.7829
24.2068 | -104.7763
24.2106 | -104.77
24.2143 | -104.763
24.218 | -104.7569
24.2218 | -104.7514
24.2255 | -104.7478
24.2292 | -104.7429
24.2329 | -104.7374
24.2367 | -104.7324
24.2404 | -104.7272
24.2441 | -104.721
24.2478 | -104.7131
24.2516 | -104.7059
24.2553 | -104.6972
24.259 | -104.6839
24.2628 | -104.6673
24.2665 | -104.6403
24.2702 | -104.6084
24.2739 | -104.5777
24.2777 | -104.5489
24.2814 | -104.5259
24.2851 | -104.5043
24.2888 | -104.4827
24.2926 | -104.4605
24.2963 | -104.4362
24.3 | -104.4088
24.3038 | -104.3811
24.3075 | -104.3519
24.3112 | -104.322
24.3149 | -104.2901
24.3187 | -104.2583
24.3224 | -104.2233
24.3261 | -104.1879
24.3298 | -104.1513
24.3336 | -104.1128
24.3373 | -104.0744
24.341 | -104.0346
24.3447 | -103.9937
24.3485 | -103.9533
24.3522 | -103.9113
24.3559 | -103.8677
24.3597 | -103.8248
24.3634 | -103.784
24.3671 | -103.7427
24.3708 | -103.7009
24.3746 | -103.6608
24.3783 | -103.6204
24.382 | -103.5799
24.3857 | -103.5393
24.3895 | -103.4998
24.3932 | -103.4624
24.3969 | -103.4265
24.4007 | -103.392
24.4044 | -103.3601
24.4081 | -103.3317
24.4118 | -103.3041
24.4156 | -103.2794
24.4193 | -103.2575
24.423 | -103.2378
24.4267 | -103.2206
24.4305 | -103.205
24.4342 | -103.1921
24.4379 | -103.1839
24.4416 | -103.1771
24.4454 | -103.1731
24.4491 | -103.1719
24.4528 | -103.1747
24.4566 | -103.179
24.4603 | -103.1863
24.464 | -103.1958
24.4677 | -103.2087
24.4715 | -103.2251
24.4752 | -103.242
24.4789 | -103.2601
24.4826 | -103.2781
24.4864 | -103.2969
24.4901 | -103.3206
24.4938 | -103.346
24.4976 | -103.3718
24.5013 | -103.3982
24.505 | -103.428
24.5087 | -103.458
24.5125 | -103.4873
24.5162 | -103.5159
24.5199 | -103.5442
24.5236 | -103.5736
24.5274 | -103.6035
24.5311 | -103.6315
24.5348 | -103.6595
24.5386 | -103.6869
24.5423 | -103.7131
24.546 | -103.7394
24.5497 | -103.7639
24.5535 | -103.7872
24.5572 | -103.8086
24.5609 | -103.8238
24.5646 | -103.8272
24.5684 | -103.8269
24.5721 | -103.8191
24.5758 | -103.8101
24.5795 | -103.7838
24.5833 | -103.7425
24.587 | -103.6927
24.5907 | -103.6158
24.5945 | -103.5206
24.5982 | -103.4246
24.6019 | -103.3334
24.6056 | -103.2652
24.6094 | -103.2002
24.6131 | -103.1359
24.6168 | -103.0651
24.6205 | -102.9838
24.6243 | -102.8946
24.628 | -102.7955
24.6317 | -102.6905
24.6355 | -102.5803
24.6392 | -102.4679
24.6429 | -102.3535
24.6466 | -102.2382
24.6504 | -102.1202
24.6541 | -102
24.6578 | -101.8787
24.6615 | -101.7574
24.6653 | -101.6356
24.669 | -101.5143
24.6727 | -101.3943
24.6764 | -101.2761
24.6802 | -101.1597
24.6839 | -101.0441
24.6876 | -100.932
24.6914 | -100.8235
24.6951 | -100.7183
24.6988 | -100.6168
24.7025 | -100.5202
24.7063 | -100.4283
24.71 | -100.3404
24.7137 | -100.257
24.7174 | -100.1799
24.7212 | -100.1079
24.7249 | -100.0423
24.7286 | -99.9837
24.7324 | -99.9789
24.7361 | -99.9348
24.7398 | -99.8969
24.7435 | -99.8653
24.7473 | -99.8402
24.751 | -99.8219
24.7547 | -99.8097
24.7584 | -99.8037
24.7622 | -99.8052
24.7659 | -99.813
24.7696 | -99.8275
24.7734 | -99.8485
24.7771 | -99.8765
24.7808 | -99.9109
24.7845 | -99.952
24.7883 | -99.9992
24.792 | -100.0533
24.7957 | -100.1133
24.7994 | -100.1794
24.8032 | -100.2506
24.8069 | -100.3279
24.8106 | -100.4102
24.8143 | -100.4982
24.8181 | -100.5902
24.8218 | -100.6861
24.8255 | -100.7871
24.8293 | -100.8925
24.833 | -101.0015
24.8367 | -101.1144
24.8404 | -101.2293
24.8442 | -101.3455
24.8479 | -101.4628
24.8516 | -101.5822
24.8553 | -101.7011
24.8591 | -101.8203
24.8628 | -101.9414
24.8665 | -102.0626
24.8703 | -102.1838
24.874 | -102.302
24.8777 | -102.4174
24.8814 | -102.5318
24.8852 | -102.6438
24.8889 | -102.7514
24.8926 | -102.8559
24.8963 | -102.9565
24.9001 | -103.0537
24.9038 | -103.146
24.9075 | -103.2352
24.9113 | -103.3228
24.915 | -103.4133
24.9187 | -103.4982
24.9224 | -103.5802
24.9262 | -103.6552
24.9299 | -103.7355
24.9336 | -103.8173
24.9373 | -103.8936
24.9411 | -103.983
24.9448 | -104.0751
24.9485 | -104.1463
24.9522 | -104.194
24.956 | -104.2046
24.9597 | -104.2033
24.9634 | -104.1994
24.9672 | -104.2
24.9709 | -104.201
24.9746 | -104.2005
24.9783 | -104.1991
24.9821 | -104.1986
24.9858 | -104.199
24.9895 | -104.1988
24.9932 | -104.1975
24.997 | -104.1943
25.0007 | -104.1903
25.0044 | -104.1876
25.0082 | -104.1856
25.0119 | -104.1828
25.0156 | -104.1798
25.0193 | -104.1786
25.0231 | -104.1762
25.0268 | -104.1719
25.0305 | -104.1687
25.0342 | -104.167
25.038 | -104.1627
25.0417 | -104.1581
25.0454 | -104.1542
25.0492 | -104.1498
25.0529 | -104.144
25.0566 | -104.1403
25.0603 | -104.1381
25.0641 | -104.1359
25.0678 | -104.1319
25.0715 | -104.1271
25.0752 | -104.1221
25.079 | -104.1157
25.0827 | -104.1106
25.0864 | -104.1078
25.0901 | -104.1038
25.0939 | -104.101
25.0976 | -104.0978
25.1013 | -104.0938
25.1051 | -104.0909
25.1088 | -104.088
25.1125 | -104.084
25.1162 | -104.0774
25.12 | -104.0723
25.1237 | -104.0686
25.1274 | -104.0655
25.1311 | -104.0617
25.1349 | -104.0591
25.1386 | -104.0579
25.1423 | -104.0562
25.1461 | -104.0539
25.1498 | -104.0523
25.1535 | -104.051
25.1572 | -104.0494
25.161 | -104.0489
25.1647 | -104.0489
25.1684 | -104.0476
25.1721 | -104.0466
25.1759 | -104.0464
25.1796 | -104.0467
25.1833 | -104.0466
25.187 | -104.0475
25.1908 | -104.047
25.1945 | -104.0466
25.1982 | -104.0486
25.202 | -104.0508
25.2057 | -104.0505
25.2094 | -104.0517
25.2131 | -104.0546
25.2169 | -104.054
25.2206 | -104.0552
25.2243 | -104.0574
25.228 | -104.0609
25.2318 | -104.0647
25.2355 | -104.0667
25.2392 | -104.0669
25.243 | -104.0651
25.2467 | -104.0635
25.2504 | -104.0642
25.2541 | -104.0661
25.2579 | -104.0691
25.2616 | -104.0707
25.2653 | -104.0718
25.269 | -104.0707
25.2728 | -104.071
25.2765 | -104.0714
25.2802 | -104.0732
25.284 | -104.0744
25.2877 | -104.0753
25.2914 | -104.0769
25.2951 | -104.078
25.2989 | -104.0778
25.3026 | -104.0776
25.3063 | -104.0789
25.31 | -104.0806
25.3138 | -104.0827
25.3175 | -104.0865
25.3212 | -104.0905
25.3249 | -104.0908
25.3287 | -104.091
25.3324 | -104.0906
25.3361 | -104.0903
25.3399 | -104.0891
25.3436 | -104.0885
25.3473 | -104.087
25.351 | -104.0875
25.3548 | -104.0882
25.3585 | -104.0887
25.3622 | -104.0904
25.3659 | -104.0898
25.3697 | -104.0879
25.3734 | -104.0873
25.3771 | -104.0864
25.3809 | -104.0867
25.3846 | -104.0892
25.3883 | -104.0918
25.392 | -104.0931
25.3958 | -104.0944
25.3995 | -104.0953
25.4032 | -104.0963
25.4069 | -104.0982
25.4107 | -104.1005
25.4144 | -104.0994
25.4181 | -104.0991
25.4218 | -104.0997
25.4256 | -104.1021
25.4293 | -104.1043
25.433 | -104.1065
25.4368 | -104.1087
25.4405 | -104.1113
25.4442 | -104.1133
25.4479 | -104.1163
25.4517 | -104.1163
25.4554 | -104.1184
25.4591 | -104.1225
25.4628 | -104.1256
25.4666 | -104.1285
25.4703 | -104.1326
25.474 | -104.1372
25.4778 | -104.1408
25.4815 | -104.1423
25.4852 | -104.1457
25.4889 | -104.1492
25.4927 | -104.1523
25.4964 | -104.155
25.5001 | -104.157
25.5038 | -104.16
25.5076 | -104.1635
25.5113 | -104.166
25.515 | -104.1697
25.5188 | -104.1725
25.5225 | -104.1736
25.5262 | -104.1756
25.5299 | -104.18
25.5337 | -104.1829
25.5374 | -104.1843
25.5411 | -104.1857
25.5448 | -104.1872
25.5486 | -104.1902
25.5523 | -104.1926
25.556 | -104.1922
25.5597 | -104.1949
25.5635 | -104.1988
25.5672 | -104.2021
25.5709 | -104.2045
25.5747 | -104.2075
25.5784 | -104.2072
25.5821 | -104.2056
25.5858 | -104.2039
25.5896 | -104.2018
25.5933 | -104.2002
25.597 | -104.1997
25.6007 | -104.2018
25.6045 | -104.2041
25.6082 | -104.2036
25.6119 | -104.2014
25.6157 | -104.1989
25.6194 | -104.1969
25.6231 | -104.1963
25.6268 | -104.1951
25.6306 | -104.1933
25.6343 | -104.1908
25.638 | -104.188
25.6417 | -104.1843
25.6455 | -104.1802
25.6492 | -104.1771
25.6529 | -104.1753
25.6566 | -104.1739
25.6604 | -104.1716
25.6641 | -104.1682
25.6678 | -104.1656
25.6716 | -104.1642
25.6753 | -104.1629
25.679 | -104.1598
25.6827 | -104.1566
25.6865 | -104.1548
25.6902 | -104.1537
25.6939 | -104.1533
25.6976 | -104.1532
25.7014 | -104.1523
25.7051 | -104.1529
25.7088 | -104.1547
25.7126 | -104.1551
25.7163 | -104.1556
25.72 | -104.1581
25.7237 | -104.1595
25.7275 | -104.1628
25.7312 | -104.168
25.7349 | -104.1726
25.7386 | -104.1776
25.7424 | -104.1823
25.7461 | -104.1865
25.7498 | -104.1913
25.7536 | -104.1985
25.7573 | -104.2073
25.761 | -104.2157
25.7647 | -104.2227
25.7685 | -104.2297
25.7722 | -104.2378
25.7759 | -104.2482
25.7796 | -104.2589
25.7834 | -104.2703
25.7871 | -104.2814
25.7908 | -104.292
25.7945 | -104.3019
25.7983 | -104.3132
25.802 | -104.3239
25.8057 | -104.3358
25.8095 | -104.3446
25.8132 | -104.3541
25.8169 | -104.3667
25.8206 | -104.377
25.8244 | -104.3843
25.8281 | -104.3945
25.8318 | -104.4053
25.8355 | -104.4166
25.8393 | -104.4282
25.843 | -104.4391
25.8467 | -104.4506
25.8505 | -104.4605
25.8542 | -104.471
25.8579 | -104.4811
25.8616 | -104.4897
25.8654 | -104.4991
25.8691 | -104.5093
25.8728 | -104.5197
25.8765 | -104.5296
25.8803 | -104.5386
25.884 | -104.5473
25.8877 | -104.5563
25.8915 | -104.5659
25.8952 | -104.5727
25.8989 | -104.5797
25.9026 | -104.5876
25.9064 | -104.5963
25.9101 | -104.6043
25.9138 | -104.6114
25.9175 | -104.6166
25.9213 | -104.6222
25.925 | -104.6258
25.9287 | -104.6295
25.9324 | -104.6329
25.9362 | -104.6373
25.9399 | -104.6419
25.9436 | -104.6468
25.9474 | -104.651
25.9511 | -104.6543
25.9548 | -104.6586
25.9585 | -104.6639
25.9623 | -104.6674
25.966 | -104.6711
25.9697 | -104.6736
25.9734 | -104.676
25.9772 | -104.6783
25.9809 | -104.681
25.9846 | -104.6828
25.9884 | -104.6824
25.9921 | -104.6846
25.9958 | -104.6876
25.9995 | -104.6902
26.0033 | -104.6933
26.007 | -104.696
26.0107 | -104.6996
26.0144 | -104.7029
26.0182 | -104.7039
26.0219 | -104.7043
26.0256 | -104.7058
26.0294 | -104.71
26.0331 | -104.7127
26.0368 | -104.715
26.0405 | -104.7184
26.0443 | -104.7226
26.048 | -104.728
26.0517 | -104.7327
26.0554 | -104.7367
26.0592 | -104.7407
26.0629 | -104.7455
26.0666 | -104.7516
26.0703 | -104.7556
26.0741 | -104.7592
26.0778 | -104.7623
26.0815 | -104.7661
26.0853 | -104.7697
26.089 | -104.7734
26.0927 | -104.7768
26.0964 | -104.781
26.1002 | -104.7848
26.1039 | -104.7876
26.1076 | -104.7907
26.1113 | -104.7939
26.1151 | -104.7979
26.1188 | -104.8013
26.1225 | -104.8037
26.1263 | -104.8077
26.13 | -104.8125
26.1337 | -104.8165
26.1374 | -104.8187
26.1412 | -104.8201
26.1449 | -104.8218
26.1486 | -104.8225
26.1523 | -104.8224
26.1561 | -104.8245
26.1598 | -104.8268
26.1635 | -104.8276
26.1672 | -104.8284
26.171 | -104.8296
26.1747 | -104.8321
26.1784 | -104.834
26.1822 | -104.8353
26.1859 | -104.8355
26.1896 | -104.8363
26.1933 | -104.8391
26.1971 | -104.8397
26.2008 | -104.8403
26.2045 | -104.8391
26.2082 | -104.8372
26.212 | -104.8367
26.2157 | -104.8366
26.2194 | -104.837
26.2232 | -104.8357
26.2269 | -104.8334
26.2306 | -104.8335
26.2343 | -104.8343
26.2381 | -104.8335
26.2418 | -104.8336
26.2455 | -104.8344
26.2492 | -104.8344
26.253 | -104.8342
26.2567 | -104.8343
26.2604 | -104.8325
26.2642 | -104.8315
26.2679 | -104.8309
26.2716 | -104.8301
26.2753 | -104.8285
26.2791 | -104.8253
26.2828 | -104.8235
26.2865 | -104.8224
26.2902 | -104.823
26.294 | -104.8225
26.2977 | -104.8215
26.3014 | -104.8203
26.3051 | -104.8174
26.3089 | -104.8151
26.3126 | -104.8137
26.3163 | -104.8126
26.3201 | -104.8131
26.3238 | -104.8124
26.3275 | -104.8112
26.3312 | -104.8101
26.335 | -104.8083
26.3387 | -104.809
26.3424 | -104.81
26.3461 | -104.8094
26.3499 | -104.8087
26.3536 | -104.8073
26.3573 | -104.8058
26.3611 | -104.806
26.3648 | -104.8062
26.3685 | -104.8068
26.3722 | -104.8081
26.376 | -104.8087
26.3797 | -104.8097
26.3834 | -104.8108
26.3871 | -104.8101
26.3909 | -104.8108
26.3946 | -104.8126
26.3983 | -104.8137
26.402 | -104.8131
26.4058 | -104.815
26.4095 | -104.8169
26.4132 | -104.818
26.417 | -104.8189
26.4207 | -104.8184
26.4244 | -104.8186
26.4281 | -104.8193
26.4319 | -104.8208
26.4356 | -104.8213
26.4393 | -104.823
26.443 | -104.8249
26.4468 | -104.8256
26.4505 | -104.8277
26.4542 | -104.8304
26.458 | -104.8321
26.4617 | -104.8338
26.4654 | -104.8358
26.4691 | -104.8347
26.4729 | -104.8332
26.4766 | -104.8327
26.4803 | -104.8335
26.484 | -104.8342
26.4878 | -104.8348
26.4915 | -104.8352
26.4952 | -104.8347
26.499 | -104.8334
26.5027 | -104.8321
26.5064 | -104.8295
26.5101 | -104.8292
26.5139 | -104.8294
26.5176 | -104.8274
26.5213 | -104.8258
26.525 | -104.825
26.5288 | -104.8236
26.5325 | -104.8231
26.5362 | -104.8221
26.5399 | -104.8218
26.5437 | -104.8215
26.5474 | -104.8197
26.5511 | -104.8181
26.5549 | -104.8161
26.5586 | -104.8157
26.5623 | -104.8151
26.566 | -104.8138
26.5698 | -104.8123
26.5735 | -104.8095
26.5772 | -104.8084
26.5809 | -104.8072
26.5847 | -104.8053
26.5884 | -104.8039
26.5921 | -104.8012
26.5959 | -104.8008
26.5996 | -104.802
26.6033 | -104.8021
26.607 | -104.8012
26.6108 | -104.8009
26.6145 | -104.8006
26.6182 | -104.7995
26.6219 | -104.799
26.6257 | -104.7998
26.6294 | -104.8016
26.6331 | -104.8031
26.6369 | -104.8028
26.6406 | -104.8034
26.6443 | -104.8044
26.648 | -104.804
26.6518 | -104.8044
26.6555 | -104.8032
26.6592 | -104.8046
26.6629 | -104.8041
26.6667 | -104.8041
26.6704 | -104.8054
26.6741 | -104.8054
26.6778 | -104.8073
26.6816 | -104.8063
26.6853 | -104.8052
26.689 | -104.8047
26.6928 | -104.803
26.6965 | -104.8015
26.7002 | -104.8007
26.7039 | -104.8006
26.7077 | -104.7998
26.7114 | -104.798
26.7151 | -104.7968
26.7188 | -104.7965
26.7226 | -104.7971
26.7263 | -104.7965
26.73 | -104.7951
26.7338 | -104.7945
26.7375 | -104.7936
26.7412 | -104.7919
26.7449 | -104.7892
26.7487 | -104.7857
26.7524 | -104.7823
26.7561 | -104.7797
26.7598 | -104.7772
26.7636 | -104.7755
26.7673 | -104.7731
26.771 | -104.7712
26.7748 | -104.7696
26.7785 | -104.7676
26.7822 | -104.7627
26.7859 | -104.7589
26.7897 | -104.7558
26.7934 | -104.7513
26.7971 | -104.7481
26.8008 | -104.7451
26.8046 | -104.744
26.8083 | -104.7412
26.812 | -104.7385
26.8157 | -104.7358
26.8195 | -104.733
26.8232 | -104.7316
26.8269 | -104.729
26.8307 | -104.7261
26.8344 | -104.7258
26.8381 | -104.7264
26.8418 | -104.726
26.8456 | -104.7269
26.8493 | -104.7276
26.853 | -104.7279
26.8567 | -104.7262
26.8605 | -104.725
26.8642 | -104.7246
26.8679 | -104.7256
26.8717 | -104.727
26.8754 | -104.7279
26.8791 | -104.7296
26.8828 | -104.7316
26.8866 | -104.7347
26.8903 | -104.7353
26.894 | -104.7348
26.8977 | -104.7335
26.9015 | -104.733
26.9052 | -104.7334
26.9089 | -104.733
26.9126 | -104.7314
26.9164 | -104.7311
26.9201 | -104.7302
26.9238 | -104.7287
26.9276 | -104.728
26.9313 | -104.7272
26.935 | -104.7268
26.9387 | -104.7254
26.9425 | -104.7229
26.9462 | -104.7188
26.9499 | -104.716
26.9536 | -104.7154
26.9574 | -104.7144
26.9611 | -104.7127
26.9648 | -104.7086
26.9686 | -104.7032
26.9723 | -104.6994
26.976 | -104.6959
26.9797 | -104.6926
26.9835 | -104.6877
26.9872 | -104.683
26.9909 | -104.6784
26.9946 | -104.6742
26.9984 | -104.6706
27.0021 | -104.6668
27.0058 | -104.6635
27.0096 | -104.6596
27.0133 | -104.6564
27.017 | -104.6539
27.0207 | -104.6519
27.0245 | -104.6462
27.0282 | -104.6429
27.0319 | -104.6413
27.0356 | -104.6395
27.0394 | -104.6365
27.0431 | -104.6329
27.0468 | -104.6287
27.0505 | -104.6244
27.0543 | -104.6232
27.058 | -104.6211
27.0617 | -104.6215
27.0655 | -104.6224
27.0692 | -104.6209
27.0729 | -104.619
27.0766 | -104.6183
27.0804 | -104.618
27.0841 | -104.6181
27.0878 | -104.618
27.0915 | -104.6159
27.0953 | -104.6142
27.099 | -104.6145
27.1027 | -104.6152
27.1065 | -104.6166
27.1102 | -104.6173
27.1139 | -104.6169
27.1176 | -104.6171
27.1214 | -104.6165
27.1251 | -104.6149
27.1288 | -104.6149
27.1325 | -104.6166
27.1363 | -104.6176
27.14 | -104.6176
27.1437 | -104.6182
27.1474 | -104.6196
27.1512 | -104.6206
27.1549 | -104.6214
27.1586 | -104.6223
27.1624 | -104.6258
27.1661 | -104.6275
27.1698 | -104.6288
27.1735 | -104.632
27.1773 | -104.6352
27.181 | -104.6376
27.1847 | -104.6363
27.1884 | -104.636
27.1922 | -104.6369
27.1959 | -104.6373
27.1996 | -104.638
27.2034 | -104.6382
27.2071 | -104.6376
27.2108 | -104.6376
27.2145 | -104.6375
27.2183 | -104.6364
27.222 | -104.6347
27.2257 | -104.6347
27.2294 | -104.6353
27.2332 | -104.6366
27.2369 | -104.6378
27.2406 | -104.6367
27.2444 | -104.6352
27.2481 | -104.6333
27.2518 | -104.6303
27.2555 | -104.6286
27.2593 | -104.6258
27.263 | -104.6228
27.2667 | -104.6205
27.2704 | -104.6179
27.2742 | -104.6165
27.2779 | -104.6142
27.2816 | -104.6124
27.2853 | -104.6096
27.2891 | -104.6058
27.2928 | -104.6017
27.2965 | -104.5965
27.3003 | -104.5911
27.304 | -104.5853
27.3077 | -104.5794
27.3114 | -104.5748
27.3152 | -104.5705
27.3189 | -104.5642
27.3226 | -104.5576
27.3263 | -104.5522
27.3301 | -104.5461
27.3338 | -104.5369
27.3375 | -104.527
27.3413 | -104.5189
27.345 | -104.5111
27.3487 | -104.5035
27.3524 | -104.4953
27.3562 | -104.4861
27.3599 | -104.4759
27.3636 | -104.4669
27.3673 | -104.4573
27.3711 | -104.4469
27.3748 | -104.4379
27.3785 | -104.4285
27.3822 | -104.4195
27.386 | -104.4105
27.3897 | -104.4026
27.3934 | -104.3941
27.3972 | -104.3861
27.4009 | -104.3786
27.4046 | -104.3701
27.4083 | -104.3652
27.4121 | -104.3576
27.4158 | -104.3507
27.4195 | -104.3468
27.4232 | -104.3392
27.427 | -104.3317
27.4307 | -104.3267
27.4344 | -104.3228
27.4382 | -104.3174
27.4419 | -104.3111
27.4456 | -104.3062
27.4493 | -104.2996
27.4531 | -104.293
27.4568 | -104.2891
27.4605 | -104.2877
27.4642 | -104.2852
27.468 | -104.2802
27.4717 | -104.2756
27.4754 | -104.2718
27.4792 | -104.269
27.4829 | -104.2652
27.4866 | -104.2613
27.4903 | -104.258
27.4941 | -104.2546
27.4978 | -104.253
27.5015 | -104.252
27.5052 | -104.2502
27.509 | -104.2479
27.5127 | -104.2462
27.5164 | -104.2445
27.5201 | -104.2416
27.5239 | -104.2392
27.5276 | -104.2383
27.5313 | -104.2362
27.5351 | -104.2334
27.5388 | -104.231
27.5425 | -104.2289
27.5462 | -104.2269
27.55 | -104.2252
27.5537 | -104.2247
27.5574 | -104.2226
27.5611 | -104.2192
27.5649 | -104.2168
27.5686 | -104.2155
27.5723 | -104.217
27.5761 | -104.2191
27.5798 | -104.2202
27.5835 | -104.2203
27.5872 | -104.2209
27.591 | -104.2224
27.5947 | -104.2241
27.5984 | -104.2258
27.6021 | -104.2272
27.6059 | -104.2288
27.6096 | -104.2311
27.6133 | -104.2339
27.6171 | -104.2361
27.6208 | -104.2369
27.6245 | -104.2388
27.6282 | -104.2433
27.632 | -104.2478
27.6357 | -104.2526
27.6394 | -104.2579
27.6431 | -104.2632
27.6469 | -104.2699
27.6506 | -104.2767
27.6543 | -104.2824
27.658 | -104.2892
27.6618 | -104.2953
27.6655 | -104.3022
27.6692 | -104.3101
27.673 | -104.317
27.6767 | -104.3235
27.6804 | -104.3308
27.6841 | -104.3386
27.6879 | -104.3464
27.6916 | -104.3548
27.6953 | -104.3622
27.699 | -104.3695
27.7028 | -104.377
27.7065 | -104.3861
27.7102 | -104.3938
27.714 | -104.4016
27.7177 | -104.4118
27.7214 | -104.4224
27.7251 | -104.4312
27.7289 | -104.4384
27.7326 | -104.4459
27.7363 | -104.4552
27.74 | -104.4648
27.7438 | -104.4729
27.7475 | -104.4806
27.7512 | -104.4878
27.755 | -104.4943
27.7587 | -104.501
27.7624 | -104.5076
27.7661 | -104.5124
27.7699 | -104.5179
27.7736 | -104.5241
27.7773 | -104.5287
27.781 | -104.5336
27.7848 | -104.5387
27.7885 | -104.5415
27.7922 | -104.5462
27.7959 | -104.5528
27.7997 | -104.5601
27.8034 | -104.5649
27.8071 | -104.5687
27.8109 | -104.5733
27.8146 | -104.5782
27.8183 | -104.5819
27.822 | -104.5856
27.8258 | -104.5901
27.8295 | -104.594
27.8332 | -104.5967
27.8369 | -104.599
27.8407 | -104.6042
27.8444 | -104.6093
27.8481 | -104.6117
27.8519 | -104.6127
27.8556 | -104.6143
27.8593 | -104.6188
27.863 | -104.6223
27.8668 | -104.6265
27.8705 | -104.6311
27.8742 | -104.6362
27.8779 | -104.6407
27.8817 | -104.6441
27.8854 | -104.6473
27.8891 | -104.6505
27.8928 | -104.6547
27.8966 | -104.6581
27.9003 | -104.6609
27.904 | -104.6648
27.9078 | -104.6687
27.9115 | -104.6719
27.9152 | -104.6759
27.9189 | -104.6773
27.9227 | -104.6784
27.9264 | -104.6786
27.9301 | -104.6716
27.9338 | -104.6636
27.9376 | -104.6588
27.9413 | -104.6558
27.945 | -104.6517
27.9488 | -104.6482
27.9525 | -104.6455
27.9562 | -104.6377
27.9599 | -104.6346
27.9637 | -104.6293
27.9674 | -104.6236
27.9711 | -104.6172
27.9748 | -104.6097
27.9786 | -104.6015
27.9823 | -104.5925
27.986 | -104.5849
27.9898 | -104.5768
27.9935 | -104.5635
27.9972 | -104.5485
28.0009 | -104.5336