## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 247 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 265
Hood | 02-92B/P | 190
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1006.4 | hPa |  | PRC PRO 01
Humidity | 55 | %/Hr
Internal temperature of calibration bench | 23.45 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1015 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.5 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.02 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.03 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3065 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 14/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 18/06/2018
Technical Manager | Anthony HUE | Approval |  | 28/06/2018
CEO | Robert GLEMOT | Approval |  | 28/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3726 | 57.3634
0.012209 | -35.9538 | 38.26
0.018313 | -34.8941 | 28.1148
0.024417 | -34.4352 | 21.8898
0.030521 | -34.3234 | 18.1302
0.036626 | -34.1785 | 15.0868
0.04273 | -34.0822 | 12.7202
0.048834 | -34.0379 | 11.193
0.054938 | -33.976 | 10.0562
0.061043 | -33.9158 | 8.9744
0.067147 | -33.9061 | 8.1574
0.073251 | -33.9074 | 7.403
0.079355 | -33.9259 | 6.7475
0.08546 | -33.94 | 6.1588
0.091564 | -33.9349 | 5.721
0.097668 | -33.9159 | 5.4724
0.10377 | -33.8934 | 5.2248
0.10988 | -33.8759 | 4.9483
0.11598 | -33.8669 | 4.6565
0.12209 | -33.8689 | 4.3731
0.12819 | -33.8792 | 4.0891
0.13429 | -33.8905 | 3.7998
0.1404 | -33.8941 | 3.521
0.1465 | -33.8963 | 3.2635
0.15261 | -33.8937 | 3.0442
0.15871 | -33.8873 | 2.8836
0.16482 | -33.8794 | 2.7826
0.17092 | -33.8774 | 2.7436
0.17702 | -33.8755 | 2.6715
0.18313 | -33.8748 | 2.5934
0.18923 | -33.8746 | 2.4851
0.19534 | -33.8736 | 2.36
0.20144 | -33.8725 | 2.2209
0.20754 | -33.8696 | 2.1042
0.21365 | -33.8687 | 1.9928
0.21975 | -33.8682 | 1.8975
0.22586 | -33.8692 | 1.8114
0.23196 | -33.8715 | 1.7389
0.23807 | -33.8755 | 1.675
0.24417 | -33.8794 | 1.6147
0.25027 | -33.8832 | 1.5512
0.25638 | -33.8855 | 1.4887
0.26248 | -33.8871 | 1.4218
0.26859 | -33.8879 | 1.3569
0.27469 | -33.8882 | 1.2864
0.2808 | -33.8894 | 1.2131
0.2869 | -33.8903 | 1.1423
0.293 | -33.8913 | 1.0755
0.29911 | -33.8924 | 1.0153
0.30521 | -33.8936 | 0.9563
0.31132 | -33.8943 | 0.91776
0.31742 | -33.8956 | 0.8752
0.32353 | -33.8959 | 0.84064
0.32963 | -33.8968 | 0.81188
0.33573 | -33.8978 | 0.79222
0.34184 | -33.8991 | 0.77142
0.34794 | -33.9012 | 0.74829
0.35405 | -33.9033 | 0.72489
0.36015 | -33.9058 | 0.70014
0.36626 | -33.9079 | 0.67326
0.37236 | -33.91 | 0.63727
0.37846 | -33.9122 | 0.59369
0.38457 | -33.9131 | 0.53701
0.39067 | -33.9142 | 0.49096
0.39678 | -33.9149 | 0.43957
0.40288 | -33.9154 | 0.38588
0.40899 | -33.9153 | 0.33428
0.41509 | -33.9147 | 0.28348
0.42119 | -33.9147 | 0.24027
0.4273 | -33.9147 | 0.19406
0.4334 | -33.9147 | 0.14786
0.43951 | -33.915 | 0.1046
0.44561 | -33.9158 | 0.064788
0.45172 | -33.9174 | 0.03298
0.45782 | -33.9182 | -0.0041062
0.46392 | -33.9194 | -0.040348
0.47003 | -33.9205 | -0.072756
0.47613 | -33.9213 | -0.10696
0.48224 | -33.9219 | -0.14174
0.48834 | -33.9209 | -0.17641
0.49445 | -33.9207 | -0.21158
0.50055 | -33.9201 | -0.24361
0.50665 | -33.9194 | -0.27766
0.51276 | -33.9184 | -0.31155
0.51886 | -33.9174 | -0.34536
0.52497 | -33.9162 | -0.36934
0.53107 | -33.9155 | -0.39859
0.53717 | -33.9148 | -0.42665
0.54328 | -33.914 | -0.44932
0.54938 | -33.9139 | -0.47177
0.55549 | -33.9139 | -0.49253
0.56159 | -33.9151 | -0.50418
0.5677 | -33.9155 | -0.51879
0.5738 | -33.9166 | -0.53515
0.5799 | -33.9175 | -0.55022
0.58601 | -33.9187 | -0.5604
0.59211 | -33.9205 | -0.56697
0.59822 | -33.9219 | -0.57664
0.60432 | -33.9235 | -0.5888
0.61043 | -33.9249 | -0.59701
0.61653 | -33.9262 | -0.60755
0.62263 | -33.9278 | -0.61712
0.62874 | -33.9285 | -0.62804
0.63484 | -33.9295 | -0.63759
0.64095 | -33.93 | -0.64804
0.64705 | -33.9305 | -0.65799
0.65316 | -33.9307 | -0.6706
0.65926 | -33.9309 | -0.68658
0.66536 | -33.9298 | -0.70391
0.67147 | -33.9294 | -0.71711
0.67757 | -33.9291 | -0.73139
0.68368 | -33.9284 | -0.75022
0.68978 | -33.9276 | -0.76731
0.69589 | -33.9267 | -0.78616
0.70199 | -33.9257 | -0.80872
0.70809 | -33.9249 | -0.82776
0.7142 | -33.9242 | -0.84806
0.7203 | -33.9231 | -0.86713
0.72641 | -33.9224 | -0.88945
0.73251 | -33.9217 | -0.90933
0.73862 | -33.9219 | -0.93106
0.74472 | -33.9216 | -0.95097
0.75082 | -33.9213 | -0.973
0.75693 | -33.9216 | -0.9948
0.76303 | -33.9222 | -1.0167
0.76914 | -33.924 | -1.0372
0.77524 | -33.9249 | -1.0568
0.78135 | -33.9262 | -1.0768
0.78745 | -33.9276 | -1.0968
0.79355 | -33.9288 | -1.1166
0.79966 | -33.9306 | -1.1361
0.80576 | -33.9329 | -1.1558
0.81187 | -33.9345 | -1.1774
0.81797 | -33.9364 | -1.1986
0.82408 | -33.9382 | -1.2191
0.83018 | -33.9399 | -1.2397
0.83628 | -33.9418 | -1.2581
0.84239 | -33.9431 | -1.2786
0.84849 | -33.9445 | -1.3003
0.8546 | -33.946 | -1.3201
0.8607 | -33.9473 | -1.3408
0.86681 | -33.9484 | -1.3597
0.87291 | -33.9495 | -1.3785
0.87901 | -33.9499 | -1.3964
0.88512 | -33.9507 | -1.4145
0.89122 | -33.9515 | -1.4327
0.89733 | -33.9519 | -1.4482
0.90343 | -33.9524 | -1.4657
0.90953 | -33.9526 | -1.4801
0.91564 | -33.9526 | -1.4959
0.92174 | -33.9529 | -1.5107
0.92785 | -33.953 | -1.5253
0.93395 | -33.9531 | -1.5391
0.94006 | -33.9533 | -1.5519
0.94616 | -33.9531 | -1.5638
0.95226 | -33.9532 | -1.5764
0.95837 | -33.9533 | -1.5902
0.96447 | -33.9536 | -1.603
0.97058 | -33.9534 | -1.6165
0.97668 | -33.9532 | -1.6298
0.98279 | -33.9529 | -1.6426
0.98889 | -33.9528 | -1.6533
0.99499 | -33.9527 | -1.6666
1.0011 | -33.9526 | -1.6773
1.0072 | -33.9523 | -1.6889
1.0133 | -33.952 | -1.6996
1.0194 | -33.9515 | -1.7125
1.0255 | -33.9512 | -1.7262
1.0316 | -33.9506 | -1.7381
1.0377 | -33.9502 | -1.7517
1.0438 | -33.9496 | -1.7644
1.0499 | -33.9489 | -1.7796
1.056 | -33.9482 | -1.7938
1.0621 | -33.9475 | -1.8081
1.0682 | -33.9471 | -1.8204
1.0743 | -33.9464 | -1.8358
1.0805 | -33.9459 | -1.8501
1.0866 | -33.9449 | -1.8651
1.0927 | -33.9441 | -1.8794
1.0988 | -33.9434 | -1.8942
1.1049 | -33.9426 | -1.9083
1.111 | -33.9418 | -1.9235
1.1171 | -33.9411 | -1.938
1.1232 | -33.9406 | -1.953
1.1293 | -33.9402 | -1.9669
1.1354 | -33.9399 | -1.9817
1.1415 | -33.9396 | -1.9971
1.1476 | -33.9389 | -2.0124
1.1537 | -33.9384 | -2.0269
1.1598 | -33.9385 | -2.0415
1.1659 | -33.9383 | -2.0546
1.172 | -33.9381 | -2.0684
1.1781 | -33.9378 | -2.0822
1.1842 | -33.9377 | -2.0962
1.1903 | -33.9381 | -2.107
1.1964 | -33.938 | -2.1203
1.2025 | -33.9381 | -2.134
1.2086 | -33.9385 | -2.1463
1.2147 | -33.9385 | -2.1586
1.2209 | -33.9389 | -2.1696
1.227 | -33.9397 | -2.1802
1.2331 | -33.9401 | -2.192
1.2392 | -33.9405 | -2.2037
1.2453 | -33.941 | -2.2146
1.2514 | -33.9414 | -2.2243
1.2575 | -33.942 | -2.2351
1.2636 | -33.9429 | -2.244
1.2697 | -33.9435 | -2.2531
1.2758 | -33.9442 | -2.2641
1.2819 | -33.945 | -2.2739
1.288 | -33.9456 | -2.2846
1.2941 | -33.9465 | -2.2929
1.3002 | -33.9472 | -2.3009
1.3063 | -33.9478 | -2.3099
1.3124 | -33.9485 | -2.3188
1.3185 | -33.9492 | -2.3278
1.3246 | -33.9499 | -2.3368
1.3307 | -33.9508 | -2.3453
1.3368 | -33.9516 | -2.3539
1.3429 | -33.9523 | -2.3628
1.349 | -33.9528 | -2.3721
1.3551 | -33.9535 | -2.3814
1.3613 | -33.9541 | -2.3896
1.3674 | -33.9547 | -2.3996
1.3735 | -33.9552 | -2.4093
1.3796 | -33.9556 | -2.4183
1.3857 | -33.956 | -2.4275
1.3918 | -33.9565 | -2.4376
1.3979 | -33.9571 | -2.4467
1.404 | -33.9576 | -2.458
1.4101 | -33.958 | -2.4683
1.4162 | -33.9584 | -2.4789
1.4223 | -33.9589 | -2.4902
1.4284 | -33.9594 | -2.5002
1.4345 | -33.9597 | -2.5117
1.4406 | -33.96 | -2.525
1.4467 | -33.9604 | -2.5366
1.4528 | -33.9609 | -2.5479
1.4589 | -33.9614 | -2.5597
1.465 | -33.9616 | -2.5705
1.4711 | -33.9621 | -2.5817
1.4772 | -33.9626 | -2.595
1.4833 | -33.9631 | -2.608
1.4894 | -33.9636 | -2.6201
1.4955 | -33.9642 | -2.632
1.5016 | -33.9647 | -2.6444
1.5078 | -33.9651 | -2.6582
1.5139 | -33.9655 | -2.6712
1.52 | -33.966 | -2.6837
1.5261 | -33.9663 | -2.6962
1.5322 | -33.9666 | -2.7101
1.5383 | -33.967 | -2.7226
1.5444 | -33.9676 | -2.7351
1.5505 | -33.968 | -2.7485
1.5566 | -33.9686 | -2.7613
1.5627 | -33.969 | -2.775
1.5688 | -33.9697 | -2.7885
1.5749 | -33.97 | -2.801
1.581 | -33.9703 | -2.8142
1.5871 | -33.9707 | -2.8278
1.5932 | -33.9711 | -2.8405
1.5993 | -33.9715 | -2.8546
1.6054 | -33.9719 | -2.8677
1.6115 | -33.9723 | -2.8817
1.6176 | -33.9726 | -2.8942
1.6237 | -33.9728 | -2.9077
1.6298 | -33.9731 | -2.9201
1.6359 | -33.9735 | -2.9327
1.642 | -33.9739 | -2.9444
1.6482 | -33.9741 | -2.9573
1.6543 | -33.9745 | -2.9708
1.6604 | -33.9748 | -2.9827
1.6665 | -33.9749 | -2.995
1.6726 | -33.9751 | -3.0074
1.6787 | -33.9755 | -3.0198
1.6848 | -33.9754 | -3.0322
1.6909 | -33.9757 | -3.0448
1.697 | -33.9759 | -3.0568
1.7031 | -33.9759 | -3.0693
1.7092 | -33.9759 | -3.0813
1.7153 | -33.9761 | -3.0931
1.7214 | -33.9759 | -3.1052
1.7275 | -33.9758 | -3.1177
1.7336 | -33.9758 | -3.1305
1.7397 | -33.9759 | -3.1427
1.7458 | -33.9759 | -3.1548
1.7519 | -33.976 | -3.1668
1.758 | -33.976 | -3.1786
1.7641 | -33.9761 | -3.1901
1.7702 | -33.9761 | -3.2029
1.7763 | -33.976 | -3.2143
1.7824 | -33.976 | -3.2266
1.7885 | -33.976 | -3.239
1.7947 | -33.9761 | -3.2509
1.8008 | -33.9761 | -3.262
1.8069 | -33.9761 | -3.2735
1.813 | -33.9762 | -3.2856
1.8191 | -33.9761 | -3.2982
1.8252 | -33.9762 | -3.3109
1.8313 | -33.9764 | -3.323
1.8374 | -33.9765 | -3.3348
1.8435 | -33.9766 | -3.3466
1.8496 | -33.9766 | -3.3597
1.8557 | -33.9768 | -3.3723
1.8618 | -33.977 | -3.3845
1.8679 | -33.9771 | -3.3965
1.874 | -33.9772 | -3.4084
1.8801 | -33.9774 | -3.4208
1.8862 | -33.9776 | -3.4326
1.8923 | -33.9777 | -3.4447
1.8984 | -33.9778 | -3.4568
1.9045 | -33.9779 | -3.4683
1.9106 | -33.9779 | -3.4805
1.9167 | -33.978 | -3.4926
1.9228 | -33.9781 | -3.5044
1.9289 | -33.9782 | -3.5163
1.9351 | -33.9784 | -3.5271
1.9412 | -33.9786 | -3.5383
1.9473 | -33.9789 | -3.5504
1.9534 | -33.9791 | -3.5618
1.9595 | -33.9792 | -3.5735
1.9656 | -33.9795 | -3.5859
1.9717 | -33.9797 | -3.5975
1.9778 | -33.98 | -3.6089
1.9839 | -33.9803 | -3.6203
1.99 | -33.9804 | -3.6314
1.9961 | -33.9805 | -3.6435
2.0022 | -33.9808 | -3.6549
2.0083 | -33.9809 | -3.6661
2.0144 | -33.9811 | -3.6775
2.0205 | -33.9814 | -3.6892
2.0266 | -33.9817 | -3.7006
2.0327 | -33.9819 | -3.7122
2.0388 | -33.9824 | -3.7237
2.0449 | -33.9825 | -3.7351
2.051 | -33.9828 | -3.7464
2.0571 | -33.9831 | -3.7578
2.0632 | -33.9833 | -3.7692
2.0693 | -33.9835 | -3.7814
2.0754 | -33.9838 | -3.7927
2.0816 | -33.9841 | -3.8048
2.0877 | -33.9843 | -3.8163
2.0938 | -33.9847 | -3.8284
2.0999 | -33.9849 | -3.84
2.106 | -33.9851 | -3.8516
2.1121 | -33.9853 | -3.8635
2.1182 | -33.9855 | -3.8743
2.1243 | -33.9858 | -3.885
2.1304 | -33.986 | -3.8961
2.1365 | -33.9863 | -3.9078
2.1426 | -33.9865 | -3.9197
2.1487 | -33.9868 | -3.9309
2.1548 | -33.9873 | -3.9427
2.1609 | -33.9877 | -3.9541
2.167 | -33.9881 | -3.9652
2.1731 | -33.9884 | -3.9772
2.1792 | -33.9886 | -3.9882
2.1853 | -33.989 | -4.0001
2.1914 | -33.9893 | -4.0115
2.1975 | -33.9897 | -4.023
2.2036 | -33.99 | -4.035
2.2097 | -33.9902 | -4.0465
2.2158 | -33.9906 | -4.0571
2.222 | -33.9909 | -4.0675
2.2281 | -33.9911 | -4.0789
2.2342 | -33.9915 | -4.0902
2.2403 | -33.9917 | -4.1008
2.2464 | -33.992 | -4.1115
2.2525 | -33.9922 | -4.1228
2.2586 | -33.9926 | -4.1334
2.2647 | -33.9929 | -4.1437
2.2708 | -33.9932 | -4.1549
2.2769 | -33.9936 | -4.1656
2.283 | -33.9938 | -4.1767
2.2891 | -33.994 | -4.1876
2.2952 | -33.9942 | -4.198
2.3013 | -33.9944 | -4.2089
2.3074 | -33.9947 | -4.2193
2.3135 | -33.9949 | -4.2306
2.3196 | -33.9952 | -4.2408
2.3257 | -33.9954 | -4.2514
2.3318 | -33.9955 | -4.2625
2.3379 | -33.9957 | -4.2732
2.344 | -33.9959 | -4.2846
2.3501 | -33.9961 | -4.2953
2.3562 | -33.9963 | -4.3056
2.3623 | -33.9964 | -4.316
2.3685 | -33.9965 | -4.3265
2.3746 | -33.9967 | -4.3376
2.3807 | -33.9969 | -4.3481
2.3868 | -33.9971 | -4.3588
2.3929 | -33.9973 | -4.3694
2.399 | -33.9973 | -4.3799
2.4051 | -33.9976 | -4.3904
2.4112 | -33.9978 | -4.401
2.4173 | -33.998 | -4.4116
2.4234 | -33.9981 | -4.4224
2.4295 | -33.9982 | -4.4332
2.4356 | -33.9984 | -4.444
2.4417 | -33.9985 | -4.4551
2.4478 | -33.9986 | -4.4659
2.4539 | -33.9988 | -4.4761
2.46 | -33.9989 | -4.4876
2.4661 | -33.9991 | -4.4986
2.4722 | -33.9993 | -4.5088
2.4783 | -33.9993 | -4.5192
2.4844 | -33.9994 | -4.5301
2.4905 | -33.9995 | -4.5411
2.4966 | -33.9997 | -4.5518
2.5027 | -33.9998 | -4.5623
2.5089 | -33.9999 | -4.5732
2.515 | -33.9999 | -4.584
2.5211 | -34 | -4.5946
2.5272 | -34.0001 | -4.6055
2.5333 | -34.0002 | -4.6165
2.5394 | -34.0004 | -4.6272
2.5455 | -34.0005 | -4.6386
2.5516 | -34.0006 | -4.6488
2.5577 | -34.0008 | -4.6603
2.5638 | -34.001 | -4.6719
2.5699 | -34.001 | -4.6833
2.576 | -34.0011 | -4.6948
2.5821 | -34.0012 | -4.706
2.5882 | -34.0013 | -4.7171
2.5943 | -34.0014 | -4.7279
2.6004 | -34.0016 | -4.7388
2.6065 | -34.0017 | -4.7492
2.6126 | -34.0018 | -4.7595
2.6187 | -34.0019 | -4.7699
2.6248 | -34.0021 | -4.7807
2.6309 | -34.0022 | -4.792
2.637 | -34.0025 | -4.8036
2.6431 | -34.0026 | -4.814
2.6492 | -34.0027 | -4.8253
2.6554 | -34.0029 | -4.8361
2.6615 | -34.003 | -4.847
2.6676 | -34.0032 | -4.8584
2.6737 | -34.0035 | -4.8694
2.6798 | -34.0038 | -4.8804
2.6859 | -34.004 | -4.8916
2.692 | -34.0041 | -4.9025
2.6981 | -34.0043 | -4.9135
2.7042 | -34.0045 | -4.9246
2.7103 | -34.0047 | -4.9352
2.7164 | -34.005 | -4.9461
2.7225 | -34.0052 | -4.9573
2.7286 | -34.0055 | -4.9683
2.7347 | -34.0057 | -4.9787
2.7408 | -34.0059 | -4.9888
2.7469 | -34.0061 | -5
2.753 | -34.0064 | -5.0105
2.7591 | -34.0067 | -5.0217
2.7652 | -34.007 | -5.0331
2.7713 | -34.0071 | -5.0438
2.7774 | -34.0074 | -5.0546
2.7835 | -34.0075 | -5.0655
2.7896 | -34.0077 | -5.0759
2.7958 | -34.008 | -5.0866
2.8019 | -34.0083 | -5.0969
2.808 | -34.0086 | -5.1076
2.8141 | -34.0089 | -5.1187
2.8202 | -34.0092 | -5.1295
2.8263 | -34.0095 | -5.1403
2.8324 | -34.0098 | -5.1516
2.8385 | -34.0102 | -5.1624
2.8446 | -34.0106 | -5.173
2.8507 | -34.011 | -5.1838
2.8568 | -34.0113 | -5.1942
2.8629 | -34.0116 | -5.2047
2.869 | -34.0119 | -5.2157
2.8751 | -34.0122 | -5.2264
2.8812 | -34.0126 | -5.2376
2.8873 | -34.0129 | -5.2485
2.8934 | -34.0132 | -5.2594
2.8995 | -34.0136 | -5.2704
2.9056 | -34.0138 | -5.2815
2.9117 | -34.0142 | -5.2925
2.9178 | -34.0146 | -5.3031
2.9239 | -34.015 | -5.3146
2.93 | -34.0154 | -5.326
2.9361 | -34.0157 | -5.3368
2.9423 | -34.016 | -5.3474
2.9484 | -34.0164 | -5.3584
2.9545 | -34.0168 | -5.3693
2.9606 | -34.0172 | -5.3806
2.9667 | -34.0176 | -5.3918
2.9728 | -34.0179 | -5.403
2.9789 | -34.0183 | -5.414
2.985 | -34.0188 | -5.4253
2.9911 | -34.0191 | -5.4365
2.9972 | -34.0196 | -5.4476
3.0033 | -34.02 | -5.459
3.0094 | -34.0204 | -5.4701
3.0155 | -34.0208 | -5.481
3.0216 | -34.0212 | -5.4922
3.0277 | -34.0215 | -5.5038
3.0338 | -34.0219 | -5.5152
3.0399 | -34.0224 | -5.5264
3.046 | -34.0227 | -5.538
3.0521 | -34.023 | -5.549
3.0582 | -34.0233 | -5.5611
3.0643 | -34.0237 | -5.5725
3.0704 | -34.024 | -5.5843
3.0765 | -34.0243 | -5.5954
3.0827 | -34.0246 | -5.6066
3.0888 | -34.0249 | -5.6178
3.0949 | -34.0253 | -5.6292
3.101 | -34.0257 | -5.6406
3.1071 | -34.026 | -5.6523
3.1132 | -34.0264 | -5.6641
3.1193 | -34.0267 | -5.6755
3.1254 | -34.0271 | -5.6868
3.1315 | -34.0275 | -5.6987
3.1376 | -34.0278 | -5.7101
3.1437 | -34.0281 | -5.7216
3.1498 | -34.0285 | -5.7337
3.1559 | -34.0288 | -5.7451
3.162 | -34.0291 | -5.7572
3.1681 | -34.0293 | -5.7691
3.1742 | -34.0297 | -5.7807
3.1803 | -34.03 | -5.7923
3.1864 | -34.0304 | -5.8036
3.1925 | -34.0307 | -5.8151
3.1986 | -34.031 | -5.8271
3.2047 | -34.0312 | -5.8392
3.2108 | -34.0315 | -5.8502
3.2169 | -34.0318 | -5.862
3.223 | -34.0321 | -5.8737
3.2292 | -34.0323 | -5.8855
3.2353 | -34.0325 | -5.8963
3.2414 | -34.0328 | -5.9078
3.2475 | -34.033 | -5.9191
3.2536 | -34.0333 | -5.9306
3.2597 | -34.0336 | -5.9415
3.2658 | -34.0338 | -5.9524
3.2719 | -34.034 | -5.9635
3.278 | -34.0342 | -5.975
3.2841 | -34.0345 | -5.9861
3.2902 | -34.0348 | -5.9976
3.2963 | -34.035 | -6.009
3.3024 | -34.0351 | -6.0201
3.3085 | -34.0354 | -6.0317
3.3146 | -34.0356 | -6.0428
3.3207 | -34.0358 | -6.0543
3.3268 | -34.0361 | -6.0657
3.3329 | -34.0363 | -6.0765
3.339 | -34.0364 | -6.0873
3.3451 | -34.0366 | -6.0985
3.3512 | -34.0368 | -6.1092
3.3573 | -34.037 | -6.1204
3.3634 | -34.0372 | -6.1315
3.3696 | -34.0374 | -6.1429
3.3757 | -34.0376 | -6.1537
3.3818 | -34.0378 | -6.1645
3.3879 | -34.038 | -6.1756
3.394 | -34.0381 | -6.1869
3.4001 | -34.0383 | -6.1978
3.4062 | -34.0384 | -6.2087
3.4123 | -34.0386 | -6.2193
3.4184 | -34.0388 | -6.2306
3.4245 | -34.039 | -6.2413
3.4306 | -34.0391 | -6.252
3.4367 | -34.0393 | -6.2625
3.4428 | -34.0394 | -6.2729
3.4489 | -34.0396 | -6.2838
3.455 | -34.0397 | -6.2947
3.4611 | -34.0399 | -6.3059
3.4672 | -34.0401 | -6.3169
3.4733 | -34.0402 | -6.328
3.4794 | -34.0404 | -6.3388
3.4855 | -34.0405 | -6.3492
3.4916 | -34.0407 | -6.3595
3.4977 | -34.0409 | -6.3708
3.5038 | -34.041 | -6.3817
3.5099 | -34.0411 | -6.3919
3.5161 | -34.0413 | -6.4023
3.5222 | -34.0414 | -6.4126
3.5283 | -34.0415 | -6.4234
3.5344 | -34.0416 | -6.4342
3.5405 | -34.0418 | -6.4447
3.5466 | -34.042 | -6.4556
3.5527 | -34.0421 | -6.4664
3.5588 | -34.0422 | -6.477
3.5649 | -34.0423 | -6.4876
3.571 | -34.0424 | -6.4982
3.5771 | -34.0425 | -6.5093
3.5832 | -34.0427 | -6.5192
3.5893 | -34.0429 | -6.5298
3.5954 | -34.043 | -6.5406
3.6015 | -34.0432 | -6.5514
3.6076 | -34.0434 | -6.562
3.6137 | -34.0436 | -6.5721
3.6198 | -34.0438 | -6.5823
3.6259 | -34.0439 | -6.5932
3.632 | -34.0441 | -6.6038
3.6381 | -34.0443 | -6.6145
3.6442 | -34.0444 | -6.6251
3.6503 | -34.0446 | -6.6352
3.6565 | -34.0449 | -6.646
3.6626 | -34.045 | -6.6566
3.6687 | -34.0452 | -6.6673
3.6748 | -34.0455 | -6.6777
3.6809 | -34.0457 | -6.6883
3.687 | -34.0459 | -6.6992
3.6931 | -34.0461 | -6.7095
3.6992 | -34.0463 | -6.7201
3.7053 | -34.0465 | -6.7305
3.7114 | -34.0467 | -6.7412
3.7175 | -34.0469 | -6.7519
3.7236 | -34.0471 | -6.7619
3.7297 | -34.0473 | -6.7727
3.7358 | -34.0476 | -6.7836
3.7419 | -34.0478 | -6.7944
3.748 | -34.048 | -6.8049
3.7541 | -34.0482 | -6.8155
3.7602 | -34.0485 | -6.826
3.7663 | -34.0487 | -6.8364
3.7724 | -34.0489 | -6.847
3.7785 | -34.0492 | -6.8574
3.7846 | -34.0494 | -6.8678
3.7907 | -34.0497 | -6.8787
3.7969 | -34.05 | -6.8894
3.803 | -34.0502 | -6.8996
3.8091 | -34.0504 | -6.9102
3.8152 | -34.0507 | -6.9206
3.8213 | -34.0509 | -6.9314
3.8274 | -34.0512 | -6.9417
3.8335 | -34.0515 | -6.9524
3.8396 | -34.0518 | -6.963
3.8457 | -34.0521 | -6.9737
3.8518 | -34.0525 | -6.9842
3.8579 | -34.0528 | -6.9948
3.864 | -34.0531 | -7.0056
3.8701 | -34.0534 | -7.0165
3.8762 | -34.0536 | -7.0269
3.8823 | -34.0539 | -7.0377
3.8884 | -34.0542 | -7.0484
3.8945 | -34.0546 | -7.0593
3.9006 | -34.0549 | -7.0703
3.9067 | -34.0552 | -7.0813
3.9128 | -34.0555 | -7.092
3.9189 | -34.0558 | -7.1025
3.925 | -34.056 | -7.1133
3.9311 | -34.0563 | -7.1234
3.9372 | -34.0566 | -7.1343
3.9434 | -34.0569 | -7.145
3.9495 | -34.0573 | -7.1558
3.9556 | -34.0576 | -7.1662
3.9617 | -34.0579 | -7.1765
3.9678 | -34.0582 | -7.1869
3.9739 | -34.0585 | -7.1975
3.98 | -34.0588 | -7.2083
3.9861 | -34.0591 | -7.2188
3.9922 | -34.0593 | -7.2295
3.9983 | -34.0596 | -7.2397
4.0044 | -34.0599 | -7.2499
4.0105 | -34.0602 | -7.2606
4.0166 | -34.0605 | -7.2713
4.0227 | -34.0609 | -7.2819
4.0288 | -34.0611 | -7.2929
4.0349 | -34.0613 | -7.3033
4.041 | -34.0616 | -7.3139
4.0471 | -34.0618 | -7.3246
4.0532 | -34.0622 | -7.3349
4.0593 | -34.0625 | -7.3457
4.0654 | -34.0628 | -7.356
4.0715 | -34.0631 | -7.3671
4.0776 | -34.0634 | -7.3776
4.0838 | -34.0637 | -7.3886
4.0899 | -34.064 | -7.399
4.096 | -34.0643 | -7.4095
4.1021 | -34.0645 | -7.4199
4.1082 | -34.0648 | -7.4302
4.1143 | -34.0651 | -7.4405
4.1204 | -34.0654 | -7.4508
4.1265 | -34.0657 | -7.461
4.1326 | -34.066 | -7.4713
4.1387 | -34.0663 | -7.4816
4.1448 | -34.0666 | -7.4923
4.1509 | -34.0668 | -7.5028
4.157 | -34.067 | -7.5137
4.1631 | -34.0673 | -7.5245
4.1692 | -34.0676 | -7.5348
4.1753 | -34.0679 | -7.5453
4.1814 | -34.0681 | -7.5554
4.1875 | -34.0684 | -7.5662
4.1936 | -34.0687 | -7.5769
4.1997 | -34.0689 | -7.587
4.2058 | -34.0692 | -7.5978
4.2119 | -34.0694 | -7.6081
4.218 | -34.0697 | -7.6183
4.2241 | -34.0699 | -7.6286
4.2303 | -34.0702 | -7.6389
4.2364 | -34.0704 | -7.6493
4.2425 | -34.0706 | -7.6598
4.2486 | -34.0709 | -7.6704
4.2547 | -34.071 | -7.6806
4.2608 | -34.0712 | -7.6911
4.2669 | -34.0714 | -7.7012
4.273 | -34.0717 | -7.7118
4.2791 | -34.0719 | -7.7222
4.2852 | -34.0722 | -7.7328
4.2913 | -34.0723 | -7.7429
4.2974 | -34.0726 | -7.7536
4.3035 | -34.0728 | -7.7643
4.3096 | -34.073 | -7.7749
4.3157 | -34.0733 | -7.7851
4.3218 | -34.0735 | -7.7954
4.3279 | -34.0737 | -7.8058
4.334 | -34.0739 | -7.8161
4.3401 | -34.0741 | -7.826
4.3462 | -34.0744 | -7.8365
4.3523 | -34.0747 | -7.8464
4.3584 | -34.0748 | -7.8564
4.3645 | -34.075 | -7.8666
4.3707 | -34.0752 | -7.8767
4.3768 | -34.0755 | -7.887
4.3829 | -34.0757 | -7.8977
4.389 | -34.0759 | -7.9079
4.3951 | -34.0761 | -7.918
4.4012 | -34.0763 | -7.9287
4.4073 | -34.0765 | -7.9388
4.4134 | -34.0767 | -7.949
4.4195 | -34.077 | -7.9592
4.4256 | -34.0773 | -7.9693
4.4317 | -34.0775 | -7.9795
4.4378 | -34.0777 | -7.9894
4.4439 | -34.0779 | -7.9994
4.45 | -34.078 | -8.0096
4.4561 | -34.0782 | -8.0199
4.4622 | -34.0784 | -8.0302
4.4683 | -34.0786 | -8.0404
4.4744 | -34.0788 | -8.0509
4.4805 | -34.079 | -8.0607
4.4866 | -34.0792 | -8.0709
4.4927 | -34.0794 | -8.0814
4.4988 | -34.0796 | -8.0916
4.5049 | -34.0798 | -8.1014
4.511 | -34.0801 | -8.1113
4.5172 | -34.0803 | -8.1214
4.5233 | -34.0805 | -8.132
4.5294 | -34.0807 | -8.1421
4.5355 | -34.0809 | -8.1523
4.5416 | -34.0811 | -8.1624
4.5477 | -34.0813 | -8.1726
4.5538 | -34.0815 | -8.1824
4.5599 | -34.0818 | -8.1925
4.566 | -34.082 | -8.2026
4.5721 | -34.0822 | -8.2128
4.5782 | -34.0824 | -8.2233
4.5843 | -34.0826 | -8.2335
4.5904 | -34.0829 | -8.2437
4.5965 | -34.0831 | -8.2538
4.6026 | -34.0833 | -8.2644
4.6087 | -34.0835 | -8.2745
4.6148 | -34.0837 | -8.285
4.6209 | -34.084 | -8.2951
4.627 | -34.0842 | -8.3055
4.6331 | -34.0843 | -8.3155
4.6392 | -34.0846 | -8.3258
4.6453 | -34.0848 | -8.3359
4.6514 | -34.085 | -8.3462
4.6576 | -34.0852 | -8.3567
4.6637 | -34.0854 | -8.3666
4.6698 | -34.0856 | -8.3766
4.6759 | -34.0858 | -8.3871
4.682 | -34.086 | -8.3972
4.6881 | -34.0861 | -8.4074
4.6942 | -34.0864 | -8.4175
4.7003 | -34.0866 | -8.4277
4.7064 | -34.0868 | -8.4377
4.7125 | -34.087 | -8.4482
4.7186 | -34.0872 | -8.4581
4.7247 | -34.0875 | -8.4684
4.7308 | -34.0877 | -8.4785
4.7369 | -34.0879 | -8.4888
4.743 | -34.0881 | -8.4993
4.7491 | -34.0883 | -8.5098
4.7552 | -34.0885 | -8.5202
4.7613 | -34.0887 | -8.5305
4.7674 | -34.0889 | -8.5405
4.7735 | -34.0891 | -8.5512
4.7796 | -34.0894 | -8.5617
4.7857 | -34.0896 | -8.572
4.7918 | -34.0898 | -8.5821
4.7979 | -34.0901 | -8.5922
4.8041 | -34.0903 | -8.6025
4.8102 | -34.0905 | -8.6124
4.8163 | -34.0908 | -8.6225
4.8224 | -34.091 | -8.6328
4.8285 | -34.0912 | -8.6427
4.8346 | -34.0915 | -8.6528
4.8407 | -34.0917 | -8.6631
4.8468 | -34.0919 | -8.6736
4.8529 | -34.0922 | -8.6841
4.859 | -34.0924 | -8.6947
4.8651 | -34.0926 | -8.7047
4.8712 | -34.0929 | -8.7147
4.8773 | -34.093 | -8.725
4.8834 | -34.0933 | -8.7353
4.8895 | -34.0935 | -8.7453
4.8956 | -34.0937 | -8.7556
4.9017 | -34.0939 | -8.7656
4.9078 | -34.0941 | -8.7762
4.9139 | -34.0944 | -8.786
4.92 | -34.0947 | -8.7959
4.9261 | -34.0949 | -8.8058
4.9322 | -34.0952 | -8.8161
4.9383 | -34.0954 | -8.8263
4.9445 | -34.0956 | -8.8366
4.9506 | -34.0959 | -8.8466
4.9567 | -34.0961 | -8.8566
4.9628 | -34.0963 | -8.8668
4.9689 | -34.0966 | -8.8774
4.975 | -34.0968 | -8.8879
4.9811 | -34.097 | -8.8979
4.9872 | -34.0973 | -8.908
4.9933 | -34.0976 | -8.9185
4.9994 | -34.0979 | -8.9291
5.0055 | -34.0981 | -8.9394
5.0116 | -34.0984 | -8.9501
5.0177 | -34.0987 | -8.9606
5.0238 | -34.099 | -8.9712
5.0299 | -34.0993 | -8.9818
5.036 | -34.0995 | -8.9921
5.0421 | -34.0998 | -9.0023
5.0482 | -34.1 | -9.0126
5.0543 | -34.1003 | -9.0227
5.0604 | -34.1006 | -9.0329
5.0665 | -34.1009 | -9.043
5.0726 | -34.1011 | -9.0535
5.0787 | -34.1014 | -9.0641
5.0848 | -34.1017 | -9.0744
5.091 | -34.1019 | -9.0845
5.0971 | -34.1022 | -9.0945
5.1032 | -34.1025 | -9.1051
5.1093 | -34.1028 | -9.1154
5.1154 | -34.103 | -9.1255
5.1215 | -34.1032 | -9.1355
5.1276 | -34.1035 | -9.1456
5.1337 | -34.1039 | -9.1556
5.1398 | -34.1041 | -9.1657
5.1459 | -34.1044 | -9.1762
5.152 | -34.1046 | -9.1867
5.1581 | -34.1049 | -9.1972
5.1642 | -34.1052 | -9.2076
5.1703 | -34.1055 | -9.2179
5.1764 | -34.1058 | -9.2281
5.1825 | -34.106 | -9.2381
5.1886 | -34.1063 | -9.2481
5.1947 | -34.1066 | -9.2584
5.2008 | -34.1069 | -9.2689
5.2069 | -34.1071 | -9.2794
5.213 | -34.1074 | -9.2895
5.2191 | -34.1077 | -9.2999
5.2252 | -34.108 | -9.3103
5.2314 | -34.1082 | -9.3208
5.2375 | -34.1085 | -9.3309
5.2436 | -34.1088 | -9.3412
5.2497 | -34.1091 | -9.3513
5.2558 | -34.1094 | -9.362
5.2619 | -34.1097 | -9.3726
5.268 | -34.11 | -9.3831
5.2741 | -34.1103 | -9.3935
5.2802 | -34.1106 | -9.4042
5.2863 | -34.1109 | -9.4145
5.2924 | -34.1112 | -9.425
5.2985 | -34.1115 | -9.4355
5.3046 | -34.1118 | -9.4459
5.3107 | -34.1121 | -9.4565
5.3168 | -34.1124 | -9.4669
5.3229 | -34.1126 | -9.4773
5.329 | -34.1129 | -9.4877
5.3351 | -34.1132 | -9.4978
5.3412 | -34.1135 | -9.5081
5.3473 | -34.1138 | -9.5186
5.3534 | -34.1141 | -9.5291
5.3595 | -34.1144 | -9.5396
5.3656 | -34.1147 | -9.55
5.3717 | -34.115 | -9.5607
5.3779 | -34.1153 | -9.5715
5.384 | -34.1156 | -9.5819
5.3901 | -34.1159 | -9.5925
5.3962 | -34.1162 | -9.6028
5.4023 | -34.1165 | -9.6134
5.4084 | -34.1169 | -9.6238
5.4145 | -34.1172 | -9.634
5.4206 | -34.1174 | -9.6444
5.4267 | -34.1178 | -9.6552
5.4328 | -34.1181 | -9.6657
5.4389 | -34.1184 | -9.6758
5.445 | -34.1186 | -9.6867
5.4511 | -34.1189 | -9.6969
5.4572 | -34.1193 | -9.7074
5.4633 | -34.1196 | -9.7178
5.4694 | -34.1198 | -9.7283
5.4755 | -34.1201 | -9.7385
5.4816 | -34.1204 | -9.749
5.4877 | -34.1208 | -9.7593
5.4938 | -34.121 | -9.7695
5.4999 | -34.1214 | -9.7798
5.506 | -34.1217 | -9.7902
5.5121 | -34.122 | -9.8006
5.5183 | -34.1223 | -9.8108
5.5244 | -34.1226 | -9.8215
5.5305 | -34.1229 | -9.8319
5.5366 | -34.1232 | -9.842
5.5427 | -34.1235 | -9.8526
5.5488 | -34.1237 | -9.8631
5.5549 | -34.124 | -9.8736
5.561 | -34.1243 | -9.8843
5.5671 | -34.1247 | -9.8945
5.5732 | -34.125 | -9.9052
5.5793 | -34.1253 | -9.916
5.5854 | -34.1256 | -9.9264
5.5915 | -34.1259 | -9.9372
5.5976 | -34.1262 | -9.9477
5.6037 | -34.1265 | -9.9582
5.6098 | -34.1268 | -9.9687
5.6159 | -34.1271 | -9.9789
5.622 | -34.1274 | -9.9895
5.6281 | -34.1277 | -10.0001
5.6342 | -34.128 | -10.0108
5.6403 | -34.1283 | -10.0213
5.6464 | -34.1286 | -10.032
5.6525 | -34.1289 | -10.0425
5.6586 | -34.1293 | -10.0527
5.6648 | -34.1296 | -10.0632
5.6709 | -34.13 | -10.0739
5.677 | -34.1303 | -10.0844
5.6831 | -34.1306 | -10.0949
5.6892 | -34.1308 | -10.1056
5.6953 | -34.1311 | -10.1161
5.7014 | -34.1315 | -10.1267
5.7075 | -34.1317 | -10.1371
5.7136 | -34.132 | -10.1475
5.7197 | -34.1323 | -10.158
5.7258 | -34.1326 | -10.1682
5.7319 | -34.1329 | -10.1787
5.738 | -34.1333 | -10.1892
5.7441 | -34.1336 | -10.1996
5.7502 | -34.1339 | -10.2098
5.7563 | -34.1342 | -10.2203
5.7624 | -34.1346 | -10.2309
5.7685 | -34.1349 | -10.2414
5.7746 | -34.1352 | -10.2522
5.7807 | -34.1354 | -10.2627
5.7868 | -34.1357 | -10.2734
5.7929 | -34.136 | -10.2838
5.799 | -34.1363 | -10.2941
5.8052 | -34.1366 | -10.3044
5.8113 | -34.1369 | -10.3148
5.8174 | -34.1372 | -10.3251
5.8235 | -34.1375 | -10.3355
5.8296 | -34.1378 | -10.3462
5.8357 | -34.1381 | -10.3565
5.8418 | -34.1384 | -10.3671
5.8479 | -34.1386 | -10.3777
5.854 | -34.139 | -10.3883
5.8601 | -34.1393 | -10.3986
5.8662 | -34.1396 | -10.4092
5.8723 | -34.1398 | -10.4195
5.8784 | -34.1401 | -10.4298
5.8845 | -34.1404 | -10.4402
5.8906 | -34.1407 | -10.4507
5.8967 | -34.141 | -10.4609
5.9028 | -34.1414 | -10.4712
5.9089 | -34.1417 | -10.4815
5.915 | -34.142 | -10.4917
5.9211 | -34.1423 | -10.5021
5.9272 | -34.1426 | -10.5122
5.9333 | -34.1429 | -10.5225
5.9394 | -34.1432 | -10.533
5.9455 | -34.1435 | -10.5431
5.9517 | -34.1439 | -10.5533
5.9578 | -34.1442 | -10.5639
5.9639 | -34.1445 | -10.5739
5.97 | -34.1449 | -10.5843
5.9761 | -34.1452 | -10.5946
5.9822 | -34.1455 | -10.6051
5.9883 | -34.1458 | -10.6154
5.9944 | -34.1461 | -10.6259
6.0005 | -34.1464 | -10.6364
6.0066 | -34.1467 | -10.6466
6.0127 | -34.1471 | -10.6567
6.0188 | -34.1474 | -10.6673
6.0249 | -34.1477 | -10.6778
6.031 | -34.148 | -10.688
6.0371 | -34.1483 | -10.6983
6.0432 | -34.1486 | -10.7086
6.0493 | -34.1489 | -10.7187
6.0554 | -34.1493 | -10.7289
6.0615 | -34.1496 | -10.7392
6.0676 | -34.1499 | -10.7498
6.0737 | -34.1502 | -10.7602
6.0798 | -34.1505 | -10.7705
6.0859 | -34.1509 | -10.7808
6.0921 | -34.1512 | -10.7909
6.0982 | -34.1515 | -10.8012
6.1043 | -34.1518 | -10.8114
6.1104 | -34.1522 | -10.8219
6.1165 | -34.1525 | -10.8317
6.1226 | -34.1528 | -10.842
6.1287 | -34.1531 | -10.8522
6.1348 | -34.1535 | -10.8625
6.1409 | -34.1538 | -10.8727
6.147 | -34.1541 | -10.8829
6.1531 | -34.1544 | -10.8931
6.1592 | -34.1547 | -10.9033
6.1653 | -34.155 | -10.9134
6.1714 | -34.1553 | -10.9237
6.1775 | -34.1556 | -10.9341
6.1836 | -34.1559 | -10.9444
6.1897 | -34.1563 | -10.9546
6.1958 | -34.1566 | -10.9651
6.2019 | -34.1569 | -10.9755
6.208 | -34.1573 | -10.9858
6.2141 | -34.1576 | -10.996
6.2202 | -34.1579 | -11.0063
6.2263 | -34.1582 | -11.0165
6.2325 | -34.1585 | -11.0264
6.2386 | -34.1588 | -11.0364
6.2447 | -34.1592 | -11.0469
6.2508 | -34.1595 | -11.0571
6.2569 | -34.1599 | -11.0673
6.263 | -34.1602 | -11.0775
6.2691 | -34.1605 | -11.0876
6.2752 | -34.1608 | -11.0979
6.2813 | -34.1611 | -11.1081
6.2874 | -34.1614 | -11.1182
6.2935 | -34.1617 | -11.1283
6.2996 | -34.162 | -11.1385
6.3057 | -34.1623 | -11.1488
6.3118 | -34.1626 | -11.1591
6.3179 | -34.1629 | -11.1692
6.324 | -34.1632 | -11.1794
6.3301 | -34.1635 | -11.1896
6.3362 | -34.1639 | -11.1997
6.3423 | -34.1642 | -11.21
6.3484 | -34.1645 | -11.2203
6.3545 | -34.1648 | -11.2301
6.3606 | -34.1651 | -11.24
6.3667 | -34.1654 | -11.25
6.3728 | -34.1657 | -11.2603
6.379 | -34.166 | -11.2704
6.3851 | -34.1664 | -11.2804
6.3912 | -34.1666 | -11.2905
6.3973 | -34.1669 | -11.3007
6.4034 | -34.1672 | -11.3106
6.4095 | -34.1675 | -11.3208
6.4156 | -34.1679 | -11.3309
6.4217 | -34.1682 | -11.341
6.4278 | -34.1685 | -11.3513
6.4339 | -34.1688 | -11.3613
6.44 | -34.1692 | -11.3714
6.4461 | -34.1694 | -11.3817
6.4522 | -34.1698 | -11.3916
6.4583 | -34.17 | -11.4019
6.4644 | -34.1703 | -11.4121
6.4705 | -34.1706 | -11.4219
6.4766 | -34.171 | -11.4321
6.4827 | -34.1713 | -11.4422
6.4888 | -34.1715 | -11.4522
6.4949 | -34.1719 | -11.462
6.501 | -34.1722 | -11.4722
6.5071 | -34.1725 | -11.4821
6.5132 | -34.1728 | -11.4922
6.5194 | -34.1731 | -11.5024
6.5255 | -34.1734 | -11.5122
6.5316 | -34.1737 | -11.5224
6.5377 | -34.174 | -11.5325
6.5438 | -34.1743 | -11.5424
6.5499 | -34.1746 | -11.5523
6.556 | -34.1749 | -11.5624
6.5621 | -34.1752 | -11.5727
6.5682 | -34.1755 | -11.5828
6.5743 | -34.1759 | -11.5928
6.5804 | -34.1762 | -11.603
6.5865 | -34.1765 | -11.6131
6.5926 | -34.1768 | -11.6231
6.5987 | -34.1771 | -11.6328
6.6048 | -34.1774 | -11.6429
6.6109 | -34.1777 | -11.653
6.617 | -34.178 | -11.6632
6.6231 | -34.1783 | -11.6732
6.6292 | -34.1786 | -11.6832
6.6353 | -34.1789 | -11.6936
6.6414 | -34.1793 | -11.7035
6.6475 | -34.1796 | -11.7136
6.6536 | -34.1798 | -11.7235
6.6597 | -34.1802 | -11.7333
6.6659 | -34.1805 | -11.7433
6.672 | -34.1808 | -11.7532
6.6781 | -34.1811 | -11.7633
6.6842 | -34.1814 | -11.7731
6.6903 | -34.1818 | -11.7831
6.6964 | -34.1821 | -11.7932
6.7025 | -34.1824 | -11.803
6.7086 | -34.1827 | -11.8131
6.7147 | -34.183 | -11.8232
6.7208 | -34.1833 | -11.8333
6.7269 | -34.1836 | -11.8432
6.733 | -34.1838 | -11.853
6.7391 | -34.1842 | -11.863
6.7452 | -34.1845 | -11.8728
6.7513 | -34.1848 | -11.8825
6.7574 | -34.1851 | -11.8926
6.7635 | -34.1854 | -11.9024
6.7696 | -34.1857 | -11.9122
6.7757 | -34.1861 | -11.9223
6.7818 | -34.1864 | -11.9322
6.7879 | -34.1867 | -11.942
6.794 | -34.187 | -11.9518
6.8001 | -34.1873 | -11.9617
6.8063 | -34.1876 | -11.9716
6.8124 | -34.1879 | -11.9817
6.8185 | -34.1882 | -11.9919
6.8246 | -34.1885 | -12.0016
6.8307 | -34.1888 | -12.0115
6.8368 | -34.1891 | -12.0217
6.8429 | -34.1895 | -12.0317
6.849 | -34.1898 | -12.0417
6.8551 | -34.1901 | -12.0515
6.8612 | -34.1904 | -12.0614
6.8673 | -34.1907 | -12.0713
6.8734 | -34.191 | -12.0814
6.8795 | -34.1913 | -12.0914
6.8856 | -34.1917 | -12.1013
6.8917 | -34.192 | -12.1112
6.8978 | -34.1922 | -12.1209
6.9039 | -34.1926 | -12.1305
6.91 | -34.1929 | -12.1405
6.9161 | -34.1932 | -12.1502
6.9222 | -34.1935 | -12.1601
6.9283 | -34.1939 | -12.1697
6.9344 | -34.1942 | -12.1797
6.9405 | -34.1945 | -12.1895
6.9466 | -34.1949 | -12.1997
6.9528 | -34.1952 | -12.2097
6.9589 | -34.1955 | -12.2198
6.965 | -34.1959 | -12.2298
6.9711 | -34.1962 | -12.2398
6.9772 | -34.1966 | -12.2497
6.9833 | -34.1968 | -12.2596
6.9894 | -34.1972 | -12.2697
6.9955 | -34.1975 | -12.2794
7.0016 | -34.1978 | -12.2892
7.0077 | -34.1981 | -12.2991
7.0138 | -34.1985 | -12.3089
7.0199 | -34.1988 | -12.3187
7.026 | -34.1991 | -12.3286
7.0321 | -34.1994 | -12.3382
7.0382 | -34.1998 | -12.3478
7.0443 | -34.2001 | -12.3578
7.0504 | -34.2004 | -12.3678
7.0565 | -34.2007 | -12.3777
7.0626 | -34.2011 | -12.3874
7.0687 | -34.2014 | -12.3968
7.0748 | -34.2017 | -12.4064
7.0809 | -34.2021 | -12.4162
7.087 | -34.2024 | -12.4261
7.0932 | -34.2027 | -12.436
7.0993 | -34.203 | -12.4459
7.1054 | -34.2033 | -12.4555
7.1115 | -34.2037 | -12.4652
7.1176 | -34.204 | -12.4749
7.1237 | -34.2043 | -12.4847
7.1298 | -34.2046 | -12.4944
7.1359 | -34.2049 | -12.5043
7.142 | -34.2053 | -12.5138
7.1481 | -34.2056 | -12.5235
7.1542 | -34.2059 | -12.5334
7.1603 | -34.2062 | -12.5431
7.1664 | -34.2066 | -12.5529
7.1725 | -34.2069 | -12.5628
7.1786 | -34.2073 | -12.5725
7.1847 | -34.2076 | -12.5824
7.1908 | -34.2079 | -12.5924
7.1969 | -34.2082 | -12.6023
7.203 | -34.2085 | -12.6119
7.2091 | -34.2089 | -12.6217
7.2152 | -34.2092 | -12.6315
7.2213 | -34.2095 | -12.6411
7.2274 | -34.2099 | -12.6508
7.2335 | -34.2102 | -12.6606
7.2397 | -34.2105 | -12.6703
7.2458 | -34.2109 | -12.6799
7.2519 | -34.2112 | -12.6897
7.258 | -34.2116 | -12.6996
7.2641 | -34.2119 | -12.7094
7.2702 | -34.2122 | -12.7193
7.2763 | -34.2126 | -12.7291
7.2824 | -34.2129 | -12.7391
7.2885 | -34.2132 | -12.7489
7.2946 | -34.2136 | -12.7587
7.3007 | -34.2139 | -12.7686
7.3068 | -34.2142 | -12.7786
7.3129 | -34.2145 | -12.7885
7.319 | -34.2149 | -12.7983
7.3251 | -34.2152 | -12.8079
7.3312 | -34.2155 | -12.8177
7.3373 | -34.2159 | -12.8273
7.3434 | -34.2162 | -12.837
7.3495 | -34.2166 | -12.8466
7.3556 | -34.2169 | -12.8565
7.3617 | -34.2173 | -12.8664
7.3678 | -34.2176 | -12.8762
7.3739 | -34.2179 | -12.8862
7.3801 | -34.2183 | -12.8962
7.3862 | -34.2186 | -12.9059
7.3923 | -34.219 | -12.9158
7.3984 | -34.2193 | -12.9253
7.4045 | -34.2197 | -12.9351
7.4106 | -34.22 | -12.9448
7.4167 | -34.2203 | -12.9547
7.4228 | -34.2207 | -12.9646
7.4289 | -34.2211 | -12.9743
7.435 | -34.2214 | -12.9839
7.4411 | -34.2218 | -12.9938
7.4472 | -34.2221 | -13.0036
7.4533 | -34.2224 | -13.0132
7.4594 | -34.2228 | -13.0232
7.4655 | -34.2231 | -13.0328
7.4716 | -34.2235 | -13.0427
7.4777 | -34.2238 | -13.0525
7.4838 | -34.2242 | -13.0621
7.4899 | -34.2245 | -13.0719
7.496 | -34.2248 | -13.0815
7.5021 | -34.2252 | -13.0912
7.5082 | -34.2255 | -13.101
7.5143 | -34.2259 | -13.1108
7.5204 | -34.2262 | -13.1207
7.5266 | -34.2266 | -13.1304
7.5327 | -34.2269 | -13.1401
7.5388 | -34.2272 | -13.1502
7.5449 | -34.2276 | -13.1601
7.551 | -34.2279 | -13.1697
7.5571 | -34.2283 | -13.1796
7.5632 | -34.2287 | -13.1895
7.5693 | -34.229 | -13.1992
7.5754 | -34.2294 | -13.2088
7.5815 | -34.2297 | -13.2186
7.5876 | -34.2301 | -13.2286
7.5937 | -34.2304 | -13.2385
7.5998 | -34.2308 | -13.2482
7.6059 | -34.2312 | -13.2581
7.612 | -34.2315 | -13.2679
7.6181 | -34.2319 | -13.2776
7.6242 | -34.2322 | -13.2877
7.6303 | -34.2326 | -13.2977
7.6364 | -34.2329 | -13.3076
7.6425 | -34.2332 | -13.3174
7.6486 | -34.2336 | -13.3272
7.6547 | -34.2339 | -13.3369
7.6608 | -34.2343 | -13.3467
7.667 | -34.2346 | -13.3566
7.6731 | -34.235 | -13.3665
7.6792 | -34.2354 | -13.3765
7.6853 | -34.2357 | -13.3864
7.6914 | -34.2361 | -13.3963
7.6975 | -34.2364 | -13.4062
7.7036 | -34.2368 | -13.4161
7.7097 | -34.2372 | -13.426
7.7158 | -34.2375 | -13.4359
7.7219 | -34.2379 | -13.4456
7.728 | -34.2382 | -13.4555
7.7341 | -34.2386 | -13.4655
7.7402 | -34.239 | -13.4754
7.7463 | -34.2393 | -13.4853
7.7524 | -34.2397 | -13.4954
7.7585 | -34.24 | -13.5053
7.7646 | -34.2404 | -13.5151
7.7707 | -34.2407 | -13.5251
7.7768 | -34.241 | -13.5351
7.7829 | -34.2414 | -13.5453
7.789 | -34.2417 | -13.5554
7.7951 | -34.2421 | -13.5652
7.8012 | -34.2424 | -13.5751
7.8073 | -34.2429 | -13.5853
7.8135 | -34.2432 | -13.5951
7.8196 | -34.2436 | -13.6052
7.8257 | -34.2439 | -13.6153
7.8318 | -34.2443 | -13.6253
7.8379 | -34.2446 | -13.6351
7.844 | -34.245 | -13.6449
7.8501 | -34.2453 | -13.6549
7.8562 | -34.2457 | -13.6651
7.8623 | -34.2461 | -13.6749
7.8684 | -34.2464 | -13.6849
7.8745 | -34.2468 | -13.6948
7.8806 | -34.2471 | -13.7049
7.8867 | -34.2475 | -13.7146
7.8928 | -34.2479 | -13.7243
7.8989 | -34.2482 | -13.7342
7.905 | -34.2486 | -13.7442
7.9111 | -34.2489 | -13.7542
7.9172 | -34.2493 | -13.7641
7.9233 | -34.2496 | -13.7741
7.9294 | -34.25 | -13.7841
7.9355 | -34.2504 | -13.7938
7.9416 | -34.2507 | -13.8037
7.9477 | -34.2511 | -13.8135
7.9539 | -34.2514 | -13.8234
7.96 | -34.2518 | -13.8334
7.9661 | -34.2521 | -13.8434
7.9722 | -34.2525 | -13.8532
7.9783 | -34.2529 | -13.8633
7.9844 | -34.2533 | -13.8732
7.9905 | -34.2536 | -13.8833
7.9966 | -34.254 | -13.8934
8.0027 | -34.2543 | -13.9033
8.0088 | -34.2547 | -13.9132
8.0149 | -34.255 | -13.9231
8.021 | -34.2554 | -13.9332
8.0271 | -34.2557 | -13.9434
8.0332 | -34.2561 | -13.9533
8.0393 | -34.2565 | -13.9634
8.0454 | -34.2568 | -13.9736
8.0515 | -34.2572 | -13.9835
8.0576 | -34.2576 | -13.9937
8.0637 | -34.2579 | -14.0036
8.0698 | -34.2583 | -14.0134
8.0759 | -34.2586 | -14.0233
8.082 | -34.259 | -14.0331
8.0881 | -34.2593 | -14.0431
8.0942 | -34.2597 | -14.0532
8.1004 | -34.2601 | -14.0632
8.1065 | -34.2605 | -14.0731
8.1126 | -34.2608 | -14.0831
8.1187 | -34.2612 | -14.0932
8.1248 | -34.2615 | -14.1029
8.1309 | -34.2619 | -14.1127
8.137 | -34.2622 | -14.1225
8.1431 | -34.2626 | -14.1327
8.1492 | -34.263 | -14.1426
8.1553 | -34.2633 | -14.1526
8.1614 | -34.2637 | -14.1627
8.1675 | -34.264 | -14.1727
8.1736 | -34.2644 | -14.1828
8.1797 | -34.2647 | -14.1928
8.1858 | -34.2651 | -14.2029
8.1919 | -34.2655 | -14.2128
8.198 | -34.2658 | -14.2226
8.2041 | -34.2662 | -14.2326
8.2102 | -34.2666 | -14.2426
8.2163 | -34.2669 | -14.2523
8.2224 | -34.2673 | -14.2623
8.2285 | -34.2676 | -14.272
8.2346 | -34.268 | -14.2819
8.2408 | -34.2684 | -14.2917
8.2469 | -34.2688 | -14.3017
8.253 | -34.2691 | -14.3115
8.2591 | -34.2695 | -14.3215
8.2652 | -34.2699 | -14.3315
8.2713 | -34.2703 | -14.3415
8.2774 | -34.2706 | -14.3516
8.2835 | -34.271 | -14.3618
8.2896 | -34.2714 | -14.3716
8.2957 | -34.2717 | -14.3816
8.3018 | -34.2721 | -14.3916
8.3079 | -34.2725 | -14.4013
8.314 | -34.2728 | -14.4113
8.3201 | -34.2732 | -14.4211
8.3262 | -34.2736 | -14.4311
8.3323 | -34.2739 | -14.441
8.3384 | -34.2743 | -14.4511
8.3445 | -34.2747 | -14.4609
8.3506 | -34.2751 | -14.471
8.3567 | -34.2755 | -14.481
8.3628 | -34.2758 | -14.4909
8.3689 | -34.2762 | -14.5007
8.375 | -34.2765 | -14.5107
8.3812 | -34.2769 | -14.5205
8.3873 | -34.2773 | -14.5305
8.3934 | -34.2776 | -14.5405
8.3995 | -34.278 | -14.5504
8.4056 | -34.2783 | -14.5603
8.4117 | -34.2787 | -14.5703
8.4178 | -34.2791 | -14.5802
8.4239 | -34.2794 | -14.59
8.43 | -34.2798 | -14.6
8.4361 | -34.2801 | -14.61
8.4422 | -34.2805 | -14.6198
8.4483 | -34.2809 | -14.6296
8.4544 | -34.2813 | -14.6395
8.4605 | -34.2816 | -14.6493
8.4666 | -34.282 | -14.6595
8.4727 | -34.2823 | -14.6695
8.4788 | -34.2827 | -14.6793
8.4849 | -34.2831 | -14.6893
8.491 | -34.2834 | -14.6992
8.4971 | -34.2838 | -14.7089
8.5032 | -34.2841 | -14.7186
8.5093 | -34.2845 | -14.7284
8.5154 | -34.2849 | -14.7384
8.5215 | -34.2853 | -14.7482
8.5277 | -34.2856 | -14.7581
8.5338 | -34.286 | -14.7679
8.5399 | -34.2863 | -14.7778
8.546 | -34.2867 | -14.7875
8.5521 | -34.2871 | -14.7974
8.5582 | -34.2874 | -14.8072
8.5643 | -34.2878 | -14.8169
8.5704 | -34.2882 | -14.8265
8.5765 | -34.2885 | -14.8365
8.5826 | -34.2889 | -14.8461
8.5887 | -34.2893 | -14.8559
8.5948 | -34.2897 | -14.8658
8.6009 | -34.2901 | -14.8755
8.607 | -34.2904 | -14.8855
8.6131 | -34.2908 | -14.8953
8.6192 | -34.2912 | -14.905
8.6253 | -34.2916 | -14.9148
8.6314 | -34.2919 | -14.9245
8.6375 | -34.2923 | -14.9345
8.6436 | -34.2926 | -14.9442
8.6497 | -34.293 | -14.9539
8.6558 | -34.2933 | -14.9637
8.6619 | -34.2937 | -14.9736
8.6681 | -34.2941 | -14.9836
8.6742 | -34.2944 | -14.9934
8.6803 | -34.2948 | -15.0032
8.6864 | -34.2952 | -15.0127
8.6925 | -34.2955 | -15.0224
8.6986 | -34.2959 | -15.0322
8.7047 | -34.2962 | -15.0422
8.7108 | -34.2966 | -15.052
8.7169 | -34.297 | -15.0618
8.723 | -34.2973 | -15.0714
8.7291 | -34.2977 | -15.081
8.7352 | -34.2981 | -15.091
8.7413 | -34.2984 | -15.1009
8.7474 | -34.2988 | -15.1105
8.7535 | -34.2992 | -15.1203
8.7596 | -34.2996 | -15.1301
8.7657 | -34.3 | -15.14
8.7718 | -34.3003 | -15.1498
8.7779 | -34.3007 | -15.1598
8.784 | -34.301 | -15.1696
8.7901 | -34.3014 | -15.1794
8.7962 | -34.3018 | -15.1889
8.8023 | -34.3022 | -15.1986
8.8084 | -34.3026 | -15.2084
8.8146 | -34.3029 | -15.2181
8.8207 | -34.3033 | -15.2279
8.8268 | -34.3037 | -15.2373
8.8329 | -34.304 | -15.247
8.839 | -34.3044 | -15.2567
8.8451 | -34.3048 | -15.2663
8.8512 | -34.3052 | -15.276
8.8573 | -34.3055 | -15.2858
8.8634 | -34.3059 | -15.2953
8.8695 | -34.3063 | -15.3052
8.8756 | -34.3066 | -15.3149
8.8817 | -34.307 | -15.3245
8.8878 | -34.3074 | -15.3345
8.8939 | -34.3077 | -15.3442
8.9 | -34.3081 | -15.3539
8.9061 | -34.3085 | -15.3636
8.9122 | -34.3089 | -15.3733
8.9183 | -34.3092 | -15.3831
8.9244 | -34.3096 | -15.3929
8.9305 | -34.31 | -15.4024
8.9366 | -34.3104 | -15.4121
8.9427 | -34.3108 | -15.4219
8.9488 | -34.3112 | -15.4316
8.955 | -34.3115 | -15.4412
8.9611 | -34.3119 | -15.451
8.9672 | -34.3123 | -15.4606
8.9733 | -34.3127 | -15.4702
8.9794 | -34.3131 | -15.4797
8.9855 | -34.3135 | -15.4895
8.9916 | -34.3139 | -15.4993
8.9977 | -34.3142 | -15.509
9.0038 | -34.3147 | -15.5185
9.0099 | -34.315 | -15.5282
9.016 | -34.3154 | -15.538
9.0221 | -34.3158 | -15.5475
9.0282 | -34.3162 | -15.557
9.0343 | -34.3166 | -15.5667
9.0404 | -34.317 | -15.5764
9.0465 | -34.3174 | -15.5863
9.0526 | -34.3178 | -15.5958
9.0587 | -34.3182 | -15.6056
9.0648 | -34.3186 | -15.6151
9.0709 | -34.3189 | -15.6248
9.077 | -34.3194 | -15.6346
9.0831 | -34.3197 | -15.6444
9.0892 | -34.3201 | -15.6541
9.0953 | -34.3205 | -15.6639
9.1015 | -34.3209 | -15.6735
9.1076 | -34.3213 | -15.683
9.1137 | -34.3217 | -15.6927
9.1198 | -34.3221 | -15.7023
9.1259 | -34.3225 | -15.7119
9.132 | -34.3229 | -15.7218
9.1381 | -34.3233 | -15.7313
9.1442 | -34.3237 | -15.7404
9.1503 | -34.3241 | -15.7502
9.1564 | -34.3245 | -15.7597
9.1625 | -34.3248 | -15.7694
9.1686 | -34.3253 | -15.7791
9.1747 | -34.3257 | -15.7887
9.1808 | -34.3261 | -15.7982
9.1869 | -34.3265 | -15.8078
9.193 | -34.3269 | -15.8175
9.1991 | -34.3273 | -15.8272
9.2052 | -34.3277 | -15.8368
9.2113 | -34.3281 | -15.8464
9.2174 | -34.3285 | -15.8561
9.2235 | -34.3289 | -15.8657
9.2296 | -34.3294 | -15.8752
9.2357 | -34.3297 | -15.8851
9.2419 | -34.3301 | -15.8946
9.248 | -34.3306 | -15.9041
9.2541 | -34.331 | -15.9138
9.2602 | -34.3314 | -15.9235
9.2663 | -34.3318 | -15.9329
9.2724 | -34.3322 | -15.9424
9.2785 | -34.3326 | -15.952
9.2846 | -34.333 | -15.9616
9.2907 | -34.3334 | -15.9711
9.2968 | -34.3338 | -15.9809
9.3029 | -34.3343 | -15.9906
9.309 | -34.3347 | -16.0003
9.3151 | -34.335 | -16.0098
9.3212 | -34.3355 | -16.0192
9.3273 | -34.3359 | -16.0287
9.3334 | -34.3363 | -16.0382
9.3395 | -34.3367 | -16.0477
9.3456 | -34.3371 | -16.0571
9.3517 | -34.3375 | -16.0667
9.3578 | -34.3379 | -16.0762
9.3639 | -34.3383 | -16.0857
9.37 | -34.3387 | -16.0954
9.3761 | -34.3391 | -16.1052
9.3822 | -34.3396 | -16.1148
9.3884 | -34.34 | -16.1241
9.3945 | -34.3404 | -16.1337
9.4006 | -34.3408 | -16.1432
9.4067 | -34.3412 | -16.1528
9.4128 | -34.3416 | -16.1625
9.4189 | -34.342 | -16.1722
9.425 | -34.3425 | -16.1817
9.4311 | -34.3429 | -16.1913
9.4372 | -34.3433 | -16.2009
9.4433 | -34.3437 | -16.2105
9.4494 | -34.3441 | -16.2198
9.4555 | -34.3446 | -16.2295
9.4616 | -34.345 | -16.239
9.4677 | -34.3454 | -16.2485
9.4738 | -34.3458 | -16.258
9.4799 | -34.3463 | -16.2676
9.486 | -34.3467 | -16.2772
9.4921 | -34.3471 | -16.2868
9.4982 | -34.3475 | -16.2962
9.5043 | -34.3479 | -16.3057
9.5104 | -34.3483 | -16.315
9.5165 | -34.3488 | -16.3245
9.5226 | -34.3492 | -16.3342
9.5288 | -34.3496 | -16.3435
9.5349 | -34.35 | -16.3532
9.541 | -34.3504 | -16.3625
9.5471 | -34.3508 | -16.3719
9.5532 | -34.3513 | -16.3815
9.5593 | -34.3517 | -16.391
9.5654 | -34.3521 | -16.4003
9.5715 | -34.3525 | -16.4098
9.5776 | -34.3529 | -16.4195
9.5837 | -34.3533 | -16.429
9.5898 | -34.3537 | -16.4385
9.5959 | -34.3541 | -16.448
9.602 | -34.3545 | -16.4573
9.6081 | -34.355 | -16.4668
9.6142 | -34.3554 | -16.4762
9.6203 | -34.3558 | -16.4857
9.6264 | -34.3562 | -16.4951
9.6325 | -34.3566 | -16.5046
9.6386 | -34.357 | -16.5139
9.6447 | -34.3574 | -16.5235
9.6508 | -34.3578 | -16.533
9.6569 | -34.3582 | -16.5424
9.663 | -34.3587 | -16.552
9.6691 | -34.3591 | -16.5614
9.6753 | -34.3595 | -16.5709
9.6814 | -34.3599 | -16.5802
9.6875 | -34.3603 | -16.5895
9.6936 | -34.3608 | -16.5988
9.6997 | -34.3612 | -16.6085
9.7058 | -34.3616 | -16.6178
9.7119 | -34.3621 | -16.6274
9.718 | -34.3625 | -16.637
9.7241 | -34.3629 | -16.6465
9.7302 | -34.3633 | -16.656
9.7363 | -34.3637 | -16.6655
9.7424 | -34.3642 | -16.6747
9.7485 | -34.3645 | -16.6841
9.7546 | -34.365 | -16.6934
9.7607 | -34.3654 | -16.703
9.7668 | -34.3658 | -16.7126
9.7729 | -34.3662 | -16.722
9.779 | -34.3667 | -16.7312
9.7851 | -34.3671 | -16.7408
9.7912 | -34.3676 | -16.7502
9.7973 | -34.368 | -16.7599
9.8034 | -34.3684 | -16.7692
9.8095 | -34.3689 | -16.7788
9.8157 | -34.3693 | -16.7883
9.8218 | -34.3697 | -16.7979
9.8279 | -34.3701 | -16.8072
9.834 | -34.3705 | -16.8168
9.8401 | -34.3709 | -16.8262
9.8462 | -34.3714 | -16.8353
9.8523 | -34.3718 | -16.8448
9.8584 | -34.3722 | -16.8543
9.8645 | -34.3727 | -16.8638
9.8706 | -34.3731 | -16.8732
9.8767 | -34.3734 | -16.8825
9.8828 | -34.3739 | -16.8919
9.8889 | -34.3743 | -16.9011
9.895 | -34.3747 | -16.9106
9.9011 | -34.3751 | -16.9201
9.9072 | -34.3756 | -16.9296
9.9133 | -34.376 | -16.939
9.9194 | -34.3764 | -16.9484
9.9255 | -34.3768 | -16.9578
9.9316 | -34.3773 | -16.9672
9.9377 | -34.3777 | -16.9766
9.9438 | -34.3781 | -16.986
9.9499 | -34.3785 | -16.9957
9.956 | -34.3789 | -17.0051
9.9622 | -34.3794 | -17.0146
9.9683 | -34.3798 | -17.0241
9.9744 | -34.3802 | -17.0337
9.9805 | -34.3806 | -17.0431
9.9866 | -34.381 | -17.0525
9.9927 | -34.3814 | -17.0618
9.9988 | -34.3818 | -17.0712
10.0049 | -34.3823 | -17.0806
10.011 | -34.3827 | -17.0902
10.0171 | -34.3831 | -17.0995
10.0232 | -34.3836 | -17.1089
10.0293 | -34.384 | -17.1184
10.0354 | -34.3845 | -17.1278
10.0415 | -34.3849 | -17.1373
10.0476 | -34.3853 | -17.1468
10.0537 | -34.3857 | -17.1563
10.0598 | -34.3861 | -17.1658
10.0659 | -34.3865 | -17.1752
10.072 | -34.3869 | -17.1845
10.0781 | -34.3874 | -17.1939
10.0842 | -34.3878 | -17.2033
10.0903 | -34.3882 | -17.2126
10.0964 | -34.3886 | -17.2218
10.1026 | -34.389 | -17.2312
10.1087 | -34.3895 | -17.2406
10.1148 | -34.3898 | -17.2501
10.1209 | -34.3903 | -17.2593
10.127 | -34.3907 | -17.2686
10.1331 | -34.3911 | -17.2778
10.1392 | -34.3915 | -17.2871
10.1453 | -34.392 | -17.2965
10.1514 | -34.3924 | -17.306
10.1575 | -34.3928 | -17.3155
10.1636 | -34.3932 | -17.3246
10.1697 | -34.3937 | -17.3341
10.1758 | -34.3941 | -17.3433
10.1819 | -34.3945 | -17.3526
10.188 | -34.3949 | -17.362
10.1941 | -34.3953 | -17.3713
10.2002 | -34.3957 | -17.3808
10.2063 | -34.3962 | -17.3901
10.2124 | -34.3966 | -17.3996
10.2185 | -34.397 | -17.409
10.2246 | -34.3974 | -17.4183
10.2307 | -34.3978 | -17.4277
10.2368 | -34.3982 | -17.4368
10.2429 | -34.3986 | -17.4461
10.2491 | -34.399 | -17.4556
10.2552 | -34.3995 | -17.465
10.2613 | -34.3999 | -17.4744
10.2674 | -34.4003 | -17.4838
10.2735 | -34.4007 | -17.4934
10.2796 | -34.4012 | -17.5028
10.2857 | -34.4016 | -17.5123
10.2918 | -34.402 | -17.5218
10.2979 | -34.4024 | -17.5312
10.304 | -34.4029 | -17.5403
10.3101 | -34.4033 | -17.5497
10.3162 | -34.4037 | -17.5592
10.3223 | -34.4041 | -17.5683
10.3284 | -34.4045 | -17.5776
10.3345 | -34.405 | -17.5868
10.3406 | -34.4054 | -17.5961
10.3467 | -34.4058 | -17.6055
10.3528 | -34.4062 | -17.6147
10.3589 | -34.4066 | -17.6241
10.365 | -34.407 | -17.6335
10.3711 | -34.4074 | -17.643
10.3772 | -34.4079 | -17.6523
10.3833 | -34.4083 | -17.6617
10.3895 | -34.4087 | -17.6711
10.3956 | -34.4091 | -17.6805
10.4017 | -34.4096 | -17.6897
10.4078 | -34.41 | -17.699
10.4139 | -34.4104 | -17.7085
10.42 | -34.4108 | -17.718
10.4261 | -34.4112 | -17.7274
10.4322 | -34.4116 | -17.7367
10.4383 | -34.412 | -17.746
10.4444 | -34.4124 | -17.7552
10.4505 | -34.4128 | -17.7646
10.4566 | -34.4133 | -17.7739
10.4627 | -34.4137 | -17.7831
10.4688 | -34.4141 | -17.7924
10.4749 | -34.4145 | -17.8017
10.481 | -34.4149 | -17.8112
10.4871 | -34.4154 | -17.8204
10.4932 | -34.4158 | -17.8296
10.4993 | -34.4162 | -17.839
10.5054 | -34.4166 | -17.8483
10.5115 | -34.4171 | -17.8577
10.5176 | -34.4175 | -17.867
10.5237 | -34.4179 | -17.8764
10.5298 | -34.4183 | -17.8856
10.536 | -34.4187 | -17.895
10.5421 | -34.4192 | -17.9043
10.5482 | -34.4196 | -17.9137
10.5543 | -34.42 | -17.9229
10.5604 | -34.4204 | -17.932
10.5665 | -34.4208 | -17.9413
10.5726 | -34.4212 | -17.9507
10.5787 | -34.4216 | -17.9599
10.5848 | -34.4221 | -17.969
10.5909 | -34.4225 | -17.9786
10.597 | -34.4229 | -17.988
10.6031 | -34.4233 | -17.9976
10.6092 | -34.4237 | -18.007
10.6153 | -34.4242 | -18.0164
10.6214 | -34.4246 | -18.0255
10.6275 | -34.425 | -18.0349
10.6336 | -34.4254 | -18.0442
10.6397 | -34.4258 | -18.0536
10.6458 | -34.4263 | -18.0627
10.6519 | -34.4267 | -18.0721
10.658 | -34.4271 | -18.0815
10.6641 | -34.4275 | -18.0909
10.6702 | -34.4279 | -18.1002
10.6764 | -34.4283 | -18.1096
10.6825 | -34.4288 | -18.1188
10.6886 | -34.4292 | -18.128
10.6947 | -34.4296 | -18.1373
10.7008 | -34.4301 | -18.1466
10.7069 | -34.4305 | -18.156
10.713 | -34.4309 | -18.1653
10.7191 | -34.4313 | -18.1746
10.7252 | -34.4317 | -18.1836
10.7313 | -34.4321 | -18.1929
10.7374 | -34.4325 | -18.2023
10.7435 | -34.4329 | -18.2117
10.7496 | -34.4333 | -18.2207
10.7557 | -34.4337 | -18.2301
10.7618 | -34.4342 | -18.2392
10.7679 | -34.4346 | -18.2485
10.774 | -34.435 | -18.2579
10.7801 | -34.4354 | -18.2672
10.7862 | -34.4358 | -18.2763
10.7923 | -34.4362 | -18.2856
10.7984 | -34.4366 | -18.2947
10.8045 | -34.4371 | -18.304
10.8106 | -34.4375 | -18.3134
10.8168 | -34.4379 | -18.3226
10.8229 | -34.4383 | -18.332
10.829 | -34.4387 | -18.3412
10.8351 | -34.4391 | -18.3503
10.8412 | -34.4395 | -18.3595
10.8473 | -34.4399 | -18.3686
10.8534 | -34.4404 | -18.3779
10.8595 | -34.4408 | -18.3873
10.8656 | -34.4412 | -18.3964
10.8717 | -34.4416 | -18.4057
10.8778 | -34.442 | -18.415
10.8839 | -34.4424 | -18.4243
10.89 | -34.4428 | -18.4336
10.8961 | -34.4433 | -18.4428
10.9022 | -34.4437 | -18.4521
10.9083 | -34.4441 | -18.4612
10.9144 | -34.4445 | -18.4705
10.9205 | -34.4449 | -18.4797
10.9266 | -34.4453 | -18.4889
10.9327 | -34.4457 | -18.498
10.9388 | -34.4462 | -18.5072
10.9449 | -34.4466 | -18.5164
10.951 | -34.447 | -18.5257
10.9571 | -34.4474 | -18.5349
10.9633 | -34.4478 | -18.5441
10.9694 | -34.4482 | -18.5535
10.9755 | -34.4486 | -18.5628
10.9816 | -34.449 | -18.5722
10.9877 | -34.4495 | -18.5812
10.9938 | -34.4499 | -18.5906
10.9999 | -34.4503 | -18.5997
11.006 | -34.4507 | -18.6086
11.0121 | -34.4511 | -18.6179
11.0182 | -34.4516 | -18.627
11.0243 | -34.452 | -18.6362
11.0304 | -34.4524 | -18.6453
11.0365 | -34.4528 | -18.6544
11.0426 | -34.4533 | -18.6636
11.0487 | -34.4537 | -18.6729
11.0548 | -34.4541 | -18.6819
11.0609 | -34.4545 | -18.6914
11.067 | -34.4549 | -18.7006
11.0731 | -34.4554 | -18.7098
11.0792 | -34.4558 | -18.719
11.0853 | -34.4562 | -18.7282
11.0914 | -34.4567 | -18.7375
11.0975 | -34.4571 | -18.7467
11.1037 | -34.4575 | -18.7558
11.1098 | -34.4579 | -18.7652
11.1159 | -34.4583 | -18.774
11.122 | -34.4587 | -18.7832
11.1281 | -34.4591 | -18.7924
11.1342 | -34.4595 | -18.8017
11.1403 | -34.46 | -18.8109
11.1464 | -34.4604 | -18.8202
11.1525 | -34.4608 | -18.8292
11.1586 | -34.4612 | -18.8386
11.1647 | -34.4616 | -18.8476
11.1708 | -34.462 | -18.8568
11.1769 | -34.4624 | -18.8661
11.183 | -34.4629 | -18.875
11.1891 | -34.4633 | -18.8842
11.1952 | -34.4637 | -18.8933
11.2013 | -34.4641 | -18.9024
11.2074 | -34.4646 | -18.9116
11.2135 | -34.465 | -18.921
11.2196 | -34.4654 | -18.9302
11.2257 | -34.4658 | -18.9393
11.2318 | -34.4663 | -18.9483
11.2379 | -34.4667 | -18.9576
11.244 | -34.4672 | -18.9669
11.2502 | -34.4676 | -18.9759
11.2563 | -34.468 | -18.9851
11.2624 | -34.4684 | -18.9941
11.2685 | -34.4689 | -19.0032
11.2746 | -34.4693 | -19.0123
11.2807 | -34.4697 | -19.0216
11.2868 | -34.4702 | -19.0309
11.2929 | -34.4706 | -19.0398
11.299 | -34.471 | -19.049
11.3051 | -34.4715 | -19.0583
11.3112 | -34.4719 | -19.0675
11.3173 | -34.4723 | -19.0764
11.3234 | -34.4728 | -19.0854
11.3295 | -34.4732 | -19.0946
11.3356 | -34.4736 | -19.1039
11.3417 | -34.474 | -19.1128
11.3478 | -34.4744 | -19.1222
11.3539 | -34.4748 | -19.1315
11.36 | -34.4753 | -19.1405
11.3661 | -34.4757 | -19.1494
11.3722 | -34.4761 | -19.1584
11.3783 | -34.4766 | -19.1677
11.3844 | -34.477 | -19.1768
11.3906 | -34.4774 | -19.1861
11.3967 | -34.4778 | -19.1951
11.4028 | -34.4783 | -19.2044
11.4089 | -34.4787 | -19.2134
11.415 | -34.4791 | -19.2224
11.4211 | -34.4795 | -19.2316
11.4272 | -34.4799 | -19.2408
11.4333 | -34.4803 | -19.2497
11.4394 | -34.4807 | -19.2587
11.4455 | -34.4812 | -19.2677
11.4516 | -34.4816 | -19.2769
11.4577 | -34.482 | -19.286
11.4638 | -34.4824 | -19.2949
11.4699 | -34.4829 | -19.3041
11.476 | -34.4833 | -19.3133
11.4821 | -34.4837 | -19.3224
11.4882 | -34.4842 | -19.3317
11.4943 | -34.4846 | -19.3408
11.5004 | -34.485 | -19.3496
11.5065 | -34.4855 | -19.3586
11.5126 | -34.4859 | -19.3675
11.5187 | -34.4863 | -19.3766
11.5248 | -34.4867 | -19.386
11.5309 | -34.4872 | -19.3949
11.5371 | -34.4876 | -19.4039
11.5432 | -34.488 | -19.4131
11.5493 | -34.4885 | -19.4223
11.5554 | -34.4889 | -19.4314
11.5615 | -34.4893 | -19.4405
11.5676 | -34.4897 | -19.4497
11.5737 | -34.4901 | -19.4587
11.5798 | -34.4906 | -19.4676
11.5859 | -34.4911 | -19.4766
11.592 | -34.4914 | -19.4857
11.5981 | -34.4919 | -19.4947
11.6042 | -34.4923 | -19.5038
11.6103 | -34.4927 | -19.5128
11.6164 | -34.4931 | -19.5218
11.6225 | -34.4936 | -19.5308
11.6286 | -34.494 | -19.5398
11.6347 | -34.4944 | -19.5489
11.6408 | -34.4948 | -19.5579
11.6469 | -34.4953 | -19.5669
11.653 | -34.4957 | -19.5759
11.6591 | -34.4961 | -19.5849
11.6652 | -34.4965 | -19.5939
11.6713 | -34.497 | -19.6028
11.6775 | -34.4974 | -19.6117
11.6836 | -34.4978 | -19.6207
11.6897 | -34.4983 | -19.6298
11.6958 | -34.4987 | -19.639
11.7019 | -34.4992 | -19.6482
11.708 | -34.4996 | -19.6571
11.7141 | -34.5001 | -19.6659
11.7202 | -34.5005 | -19.6748
11.7263 | -34.5009 | -19.684
11.7324 | -34.5013 | -19.6931
11.7385 | -34.5017 | -19.7023
11.7446 | -34.5021 | -19.7113
11.7507 | -34.5026 | -19.72
11.7568 | -34.503 | -19.729
11.7629 | -34.5034 | -19.738
11.769 | -34.5039 | -19.7469
11.7751 | -34.5043 | -19.7558
11.7812 | -34.5048 | -19.7646
11.7873 | -34.5052 | -19.7736
11.7934 | -34.5056 | -19.7827
11.7995 | -34.506 | -19.7916
11.8056 | -34.5065 | -19.8004
11.8117 | -34.5069 | -19.8094
11.8178 | -34.5074 | -19.8182
11.824 | -34.5078 | -19.8271
11.8301 | -34.5082 | -19.836
11.8362 | -34.5087 | -19.8449
11.8423 | -34.5091 | -19.8539
11.8484 | -34.5095 | -19.8629
11.8545 | -34.51 | -19.8717
11.8606 | -34.5104 | -19.8805
11.8667 | -34.5108 | -19.8894
11.8728 | -34.5112 | -19.8981
11.8789 | -34.5116 | -19.9071
11.885 | -34.5121 | -19.9161
11.8911 | -34.5125 | -19.9249
11.8972 | -34.513 | -19.9338
11.9033 | -34.5134 | -19.9427
11.9094 | -34.5138 | -19.9515
11.9155 | -34.5143 | -19.9604
11.9216 | -34.5147 | -19.9691
11.9277 | -34.5151 | -19.9779
11.9338 | -34.5156 | -19.9868
11.9399 | -34.516 | -19.9957
11.946 | -34.5164 | -20.0046
11.9521 | -34.5169 | -20.0134
11.9582 | -34.5173 | -20.0221
11.9644 | -34.5177 | -20.0308
11.9705 | -34.5182 | -20.0397
11.9766 | -34.5187 | -20.0486
11.9827 | -34.5191 | -20.0575
11.9888 | -34.5196 | -20.0663
11.9949 | -34.52 | -20.0748
12.001 | -34.5205 | -20.0838
12.0071 | -34.5209 | -20.0926
12.0132 | -34.5213 | -20.1012
12.0193 | -34.5217 | -20.1102
12.0254 | -34.5222 | -20.119
12.0315 | -34.5226 | -20.1275
12.0376 | -34.5231 | -20.1362
12.0437 | -34.5235 | -20.145
12.0498 | -34.5239 | -20.1539
12.0559 | -34.5244 | -20.1626
12.062 | -34.5248 | -20.1711
12.0681 | -34.5253 | -20.18
12.0742 | -34.5257 | -20.1887
12.0803 | -34.5261 | -20.1977
12.0864 | -34.5266 | -20.2065
12.0925 | -34.527 | -20.2153
12.0986 | -34.5275 | -20.224
12.1047 | -34.5279 | -20.2329
12.1109 | -34.5284 | -20.2415
12.117 | -34.5288 | -20.2503
12.1231 | -34.5293 | -20.259
12.1292 | -34.5297 | -20.2677
12.1353 | -34.5301 | -20.2764
12.1414 | -34.5306 | -20.2852
12.1475 | -34.531 | -20.2941
12.1536 | -34.5314 | -20.3029
12.1597 | -34.5319 | -20.3117
12.1658 | -34.5324 | -20.3204
12.1719 | -34.5328 | -20.3289
12.178 | -34.5332 | -20.3376
12.1841 | -34.5337 | -20.3465
12.1902 | -34.5341 | -20.3553
12.1963 | -34.5346 | -20.3642
12.2024 | -34.5351 | -20.3726
12.2085 | -34.5355 | -20.3814
12.2146 | -34.5359 | -20.3902
12.2207 | -34.5364 | -20.3989
12.2268 | -34.5368 | -20.4076
12.2329 | -34.5373 | -20.4163
12.239 | -34.5378 | -20.4246
12.2451 | -34.5382 | -20.4334
12.2513 | -34.5387 | -20.4423
12.2574 | -34.5391 | -20.4509
12.2635 | -34.5396 | -20.4596
12.2696 | -34.54 | -20.4683
12.2757 | -34.5405 | -20.4768
12.2818 | -34.5409 | -20.4856
12.2879 | -34.5414 | -20.4943
12.294 | -34.5418 | -20.5029
12.3001 | -34.5423 | -20.5117
12.3062 | -34.5428 | -20.5204
12.3123 | -34.5432 | -20.5287
12.3184 | -34.5437 | -20.5373
12.3245 | -34.5441 | -20.5459
12.3306 | -34.5446 | -20.5545
12.3367 | -34.545 | -20.5633
12.3428 | -34.5455 | -20.572
12.3489 | -34.5459 | -20.5804
12.355 | -34.5464 | -20.5892
12.3611 | -34.5468 | -20.5979
12.3672 | -34.5473 | -20.6066
12.3733 | -34.5477 | -20.6153
12.3794 | -34.5482 | -20.6236
12.3855 | -34.5486 | -20.6322
12.3916 | -34.5491 | -20.6409
12.3978 | -34.5495 | -20.6497
12.4039 | -34.55 | -20.6584
12.41 | -34.5504 | -20.6671
12.4161 | -34.5509 | -20.6753
12.4222 | -34.5514 | -20.6839
12.4283 | -34.5518 | -20.6925
12.4344 | -34.5523 | -20.7012
12.4405 | -34.5527 | -20.7099
12.4466 | -34.5532 | -20.7184
12.4527 | -34.5536 | -20.7269
12.4588 | -34.5541 | -20.7356
12.4649 | -34.5545 | -20.7443
12.471 | -34.555 | -20.753
12.4771 | -34.5554 | -20.7616
12.4832 | -34.5559 | -20.7701
12.4893 | -34.5564 | -20.7785
12.4954 | -34.5568 | -20.7871
12.5015 | -34.5573 | -20.7956
12.5076 | -34.5577 | -20.8043
12.5137 | -34.5582 | -20.813
12.5198 | -34.5587 | -20.8214
12.5259 | -34.5591 | -20.83
12.532 | -34.5596 | -20.8385
12.5382 | -34.56 | -20.8471
12.5443 | -34.5605 | -20.8556
12.5504 | -34.5609 | -20.8643
12.5565 | -34.5614 | -20.8728
12.5626 | -34.5619 | -20.8815
12.5687 | -34.5623 | -20.8902
12.5748 | -34.5628 | -20.8989
12.5809 | -34.5632 | -20.9074
12.587 | -34.5637 | -20.9161
12.5931 | -34.5642 | -20.9244
12.5992 | -34.5647 | -20.9329
12.6053 | -34.5651 | -20.9415
12.6114 | -34.5656 | -20.9501
12.6175 | -34.5661 | -20.9587
12.6236 | -34.5665 | -20.9674
12.6297 | -34.567 | -20.9757
12.6358 | -34.5674 | -20.9842
12.6419 | -34.5679 | -20.9927
12.648 | -34.5684 | -21.0013
12.6541 | -34.5688 | -21.0098
12.6602 | -34.5693 | -21.0181
12.6663 | -34.5698 | -21.0267
12.6724 | -34.5703 | -21.0351
12.6785 | -34.5707 | -21.0436
12.6847 | -34.5712 | -21.0521
12.6908 | -34.5717 | -21.0606
12.6969 | -34.5722 | -21.0691
12.703 | -34.5726 | -21.0778
12.7091 | -34.5731 | -21.0863
12.7152 | -34.5735 | -21.0946
12.7213 | -34.574 | -21.1033
12.7274 | -34.5744 | -21.1118
12.7335 | -34.5749 | -21.1201
12.7396 | -34.5754 | -21.1284
12.7457 | -34.5758 | -21.1371
12.7518 | -34.5763 | -21.1456
12.7579 | -34.5768 | -21.1542
12.764 | -34.5772 | -21.1628
12.7701 | -34.5777 | -21.1712
12.7762 | -34.5781 | -21.1796
12.7823 | -34.5786 | -21.1881
12.7884 | -34.5791 | -21.1967
12.7945 | -34.5796 | -21.2054
12.8006 | -34.5801 | -21.2139
12.8067 | -34.5805 | -21.2223
12.8128 | -34.581 | -21.2308
12.8189 | -34.5815 | -21.2392
12.8251 | -34.5819 | -21.2478
12.8312 | -34.5824 | -21.2563
12.8373 | -34.5829 | -21.2647
12.8434 | -34.5833 | -21.2733
12.8495 | -34.5838 | -21.2817
12.8556 | -34.5842 | -21.2901
12.8617 | -34.5847 | -21.2984
12.8678 | -34.5852 | -21.3071
12.8739 | -34.5856 | -21.3154
12.88 | -34.5861 | -21.3239
12.8861 | -34.5865 | -21.3323
12.8922 | -34.587 | -21.3409
12.8983 | -34.5875 | -21.3493
12.9044 | -34.5879 | -21.3576
12.9105 | -34.5884 | -21.3659
12.9166 | -34.5889 | -21.3743
12.9227 | -34.5893 | -21.3827
12.9288 | -34.5898 | -21.3913
12.9349 | -34.5902 | -21.3999
12.941 | -34.5908 | -21.4082
12.9471 | -34.5912 | -21.4167
12.9532 | -34.5917 | -21.4251
12.9593 | -34.5922 | -21.4336
12.9654 | -34.5926 | -21.4422
12.9716 | -34.5931 | -21.4507
12.9777 | -34.5936 | -21.4592
12.9838 | -34.5941 | -21.4675
12.9899 | -34.5946 | -21.476
12.996 | -34.595 | -21.4846
13.0021 | -34.5955 | -21.4932
13.0082 | -34.596 | -21.5015
13.0143 | -34.5965 | -21.5098
13.0204 | -34.5969 | -21.5182
13.0265 | -34.5974 | -21.5265
13.0326 | -34.5979 | -21.5349
13.0387 | -34.5983 | -21.5434
13.0448 | -34.5988 | -21.5518
13.0509 | -34.5992 | -21.5602
13.057 | -34.5997 | -21.5686
13.0631 | -34.6002 | -21.577
13.0692 | -34.6007 | -21.5854
13.0753 | -34.6012 | -21.594
13.0814 | -34.6016 | -21.6026
13.0875 | -34.6021 | -21.6108
13.0936 | -34.6026 | -21.6193
13.0997 | -34.603 | -21.6277
13.1058 | -34.6035 | -21.636
13.112 | -34.604 | -21.6444
13.1181 | -34.6045 | -21.6527
13.1242 | -34.6049 | -21.6611
13.1303 | -34.6054 | -21.6696
13.1364 | -34.6059 | -21.6781
13.1425 | -34.6063 | -21.6864
13.1486 | -34.6068 | -21.6949
13.1547 | -34.6073 | -21.7033
13.1608 | -34.6077 | -21.7117
13.1669 | -34.6082 | -21.7201
13.173 | -34.6087 | -21.7286
13.1791 | -34.6092 | -21.7368
13.1852 | -34.6096 | -21.7452
13.1913 | -34.6101 | -21.7534
13.1974 | -34.6105 | -21.7619
13.2035 | -34.611 | -21.7702
13.2096 | -34.6115 | -21.7786
13.2157 | -34.612 | -21.7869
13.2218 | -34.6124 | -21.7952
13.2279 | -34.6129 | -21.8034
13.234 | -34.6134 | -21.8118
13.2401 | -34.6139 | -21.8202
13.2462 | -34.6143 | -21.8286
13.2524 | -34.6148 | -21.8371
13.2585 | -34.6153 | -21.8455
13.2646 | -34.6157 | -21.8539
13.2707 | -34.6162 | -21.8623
13.2768 | -34.6167 | -21.8707
13.2829 | -34.6172 | -21.8792
13.289 | -34.6177 | -21.8875
13.2951 | -34.6181 | -21.8958
13.3012 | -34.6186 | -21.9042
13.3073 | -34.6191 | -21.9125
13.3134 | -34.6195 | -21.9209
13.3195 | -34.62 | -21.9292
13.3256 | -34.6205 | -21.9375
13.3317 | -34.6209 | -21.9459
13.3378 | -34.6214 | -21.9544
13.3439 | -34.6219 | -21.963
13.35 | -34.6224 | -21.9714
13.3561 | -34.6228 | -21.9797
13.3622 | -34.6233 | -21.9881
13.3683 | -34.6238 | -21.9964
13.3744 | -34.6242 | -22.0046
13.3805 | -34.6247 | -22.0131
13.3866 | -34.6252 | -22.0215
13.3927 | -34.6256 | -22.0298
13.3989 | -34.6261 | -22.038
13.405 | -34.6266 | -22.0463
13.4111 | -34.627 | -22.0545
13.4172 | -34.6275 | -22.0629
13.4233 | -34.628 | -22.0715
13.4294 | -34.6285 | -22.0801
13.4355 | -34.6289 | -22.0882
13.4416 | -34.6294 | -22.0965
13.4477 | -34.6299 | -22.1049
13.4538 | -34.6304 | -22.1133
13.4599 | -34.6309 | -22.1217
13.466 | -34.6313 | -22.1301
13.4721 | -34.6318 | -22.1383
13.4782 | -34.6323 | -22.1466
13.4843 | -34.6328 | -22.155
13.4904 | -34.6333 | -22.1633
13.4965 | -34.6337 | -22.1718
13.5026 | -34.6342 | -22.1803
13.5087 | -34.6347 | -22.1885
13.5148 | -34.6351 | -22.197
13.5209 | -34.6356 | -22.2055
13.527 | -34.636 | -22.2138
13.5331 | -34.6365 | -22.2222
13.5393 | -34.637 | -22.2304
13.5454 | -34.6375 | -22.2386
13.5515 | -34.638 | -22.2471
13.5576 | -34.6385 | -22.2554
13.5637 | -34.6389 | -22.2639
13.5698 | -34.6394 | -22.2723
13.5759 | -34.6399 | -22.2803
13.582 | -34.6403 | -22.2886
13.5881 | -34.6408 | -22.297
13.5942 | -34.6413 | -22.3056
13.6003 | -34.6418 | -22.3139
13.6064 | -34.6423 | -22.3224
13.6125 | -34.6427 | -22.3308
13.6186 | -34.6432 | -22.3391
13.6247 | -34.6437 | -22.3475
13.6308 | -34.6442 | -22.356
13.6369 | -34.6446 | -22.3644
13.643 | -34.6451 | -22.3728
13.6491 | -34.6456 | -22.3812
13.6552 | -34.6461 | -22.3895
13.6613 | -34.6465 | -22.3978
13.6674 | -34.647 | -22.4062
13.6735 | -34.6475 | -22.4146
13.6796 | -34.648 | -22.4229
13.6858 | -34.6484 | -22.4312
13.6919 | -34.6489 | -22.4395
13.698 | -34.6494 | -22.4479
13.7041 | -34.6499 | -22.4564
13.7102 | -34.6504 | -22.4648
13.7163 | -34.6509 | -22.4732
13.7224 | -34.6514 | -22.4814
13.7285 | -34.6519 | -22.4898
13.7346 | -34.6523 | -22.4981
13.7407 | -34.6528 | -22.5064
13.7468 | -34.6533 | -22.5149
13.7529 | -34.6538 | -22.5233
13.759 | -34.6543 | -22.5317
13.7651 | -34.6547 | -22.5402
13.7712 | -34.6552 | -22.5486
13.7773 | -34.6556 | -22.557
13.7834 | -34.6561 | -22.5652
13.7895 | -34.6566 | -22.5735
13.7956 | -34.6571 | -22.5821
13.8017 | -34.6575 | -22.5903
13.8078 | -34.658 | -22.5986
13.8139 | -34.6585 | -22.6069
13.82 | -34.6589 | -22.6151
13.8262 | -34.6594 | -22.6234
13.8323 | -34.6599 | -22.6318
13.8384 | -34.6604 | -22.6402
13.8445 | -34.6609 | -22.6486
13.8506 | -34.6614 | -22.657
13.8567 | -34.6618 | -22.6654
13.8628 | -34.6623 | -22.6738
13.8689 | -34.6628 | -22.6824
13.875 | -34.6632 | -22.6907
13.8811 | -34.6637 | -22.699
13.8872 | -34.6641 | -22.7073
13.8933 | -34.6646 | -22.7155
13.8994 | -34.665 | -22.7239
13.9055 | -34.6655 | -22.7323
13.9116 | -34.666 | -22.7407
13.9177 | -34.6665 | -22.749
13.9238 | -34.6669 | -22.7573
13.9299 | -34.6674 | -22.7655
13.936 | -34.6679 | -22.7738
13.9421 | -34.6683 | -22.7821
13.9482 | -34.6688 | -22.7904
13.9543 | -34.6692 | -22.7989
13.9604 | -34.6697 | -22.8072
13.9665 | -34.6702 | -22.8154
13.9727 | -34.6707 | -22.8239
13.9788 | -34.6712 | -22.8321
13.9849 | -34.6716 | -22.8406
13.991 | -34.6721 | -22.8489
13.9971 | -34.6726 | -22.8572
14.0032 | -34.673 | -22.8656
14.0093 | -34.6735 | -22.874
14.0154 | -34.6739 | -22.8822
14.0215 | -34.6744 | -22.8905
14.0276 | -34.6749 | -22.8988
14.0337 | -34.6754 | -22.9071
14.0398 | -34.6759 | -22.9154
14.0459 | -34.6764 | -22.9237
14.052 | -34.6768 | -22.9322
14.0581 | -34.6773 | -22.9406
14.0642 | -34.6778 | -22.9489
14.0703 | -34.6783 | -22.9573
14.0764 | -34.6787 | -22.9655
14.0825 | -34.6792 | -22.9737
14.0886 | -34.6797 | -22.9821
14.0947 | -34.6801 | -22.9905
14.1008 | -34.6806 | -22.9987
14.1069 | -34.681 | -23.0069
14.1131 | -34.6815 | -23.0153
14.1192 | -34.682 | -23.0235
14.1253 | -34.6824 | -23.0317
14.1314 | -34.6829 | -23.04
14.1375 | -34.6834 | -23.0482
14.1436 | -34.6839 | -23.0565
14.1497 | -34.6844 | -23.0647
14.1558 | -34.6848 | -23.073
14.1619 | -34.6853 | -23.0813
14.168 | -34.6858 | -23.0895
14.1741 | -34.6863 | -23.0978
14.1802 | -34.6867 | -23.106
14.1863 | -34.6872 | -23.1143
14.1924 | -34.6877 | -23.1226
14.1985 | -34.6882 | -23.1309
14.2046 | -34.6887 | -23.1393
14.2107 | -34.6892 | -23.1475
14.2168 | -34.6896 | -23.1559
14.2229 | -34.6901 | -23.1642
14.229 | -34.6906 | -23.1726
14.2351 | -34.691 | -23.1809
14.2412 | -34.6915 | -23.1892
14.2473 | -34.692 | -23.1974
14.2534 | -34.6925 | -23.2056
14.2596 | -34.6929 | -23.2137
14.2657 | -34.6933 | -23.222
14.2718 | -34.6938 | -23.2303
14.2779 | -34.6943 | -23.2387
14.284 | -34.6948 | -23.247
14.2901 | -34.6952 | -23.2552
14.2962 | -34.6957 | -23.2636
14.3023 | -34.6962 | -23.2717
14.3084 | -34.6967 | -23.28
14.3145 | -34.6972 | -23.2883
14.3206 | -34.6976 | -23.2968
14.3267 | -34.6981 | -23.3051
14.3328 | -34.6986 | -23.3133
14.3389 | -34.699 | -23.3216
14.345 | -34.6995 | -23.3299
14.3511 | -34.7 | -23.3382
14.3572 | -34.7004 | -23.3467
14.3633 | -34.7009 | -23.355
14.3694 | -34.7014 | -23.3633
14.3755 | -34.7019 | -23.3717
14.3816 | -34.7024 | -23.3801
14.3877 | -34.7028 | -23.3883
14.3938 | -34.7033 | -23.3967
14.4 | -34.7038 | -23.405
14.4061 | -34.7043 | -23.4133
14.4122 | -34.7047 | -23.4216
14.4183 | -34.7052 | -23.4298
14.4244 | -34.7057 | -23.4381
14.4305 | -34.7062 | -23.4463
14.4366 | -34.7067 | -23.4546
14.4427 | -34.7071 | -23.463
14.4488 | -34.7076 | -23.4714
14.4549 | -34.7081 | -23.4797
14.461 | -34.7085 | -23.488
14.4671 | -34.709 | -23.4964
14.4732 | -34.7095 | -23.5048
14.4793 | -34.71 | -23.5131
14.4854 | -34.7104 | -23.5213
14.4915 | -34.7109 | -23.5293
14.4976 | -34.7114 | -23.5375
14.5037 | -34.7119 | -23.5457
14.5098 | -34.7123 | -23.5539
14.5159 | -34.7128 | -23.5622
14.522 | -34.7133 | -23.5703
14.5281 | -34.7138 | -23.5784
14.5342 | -34.7142 | -23.5868
14.5403 | -34.7147 | -23.595
14.5465 | -34.7152 | -23.6033
14.5526 | -34.7157 | -23.6118
14.5587 | -34.7162 | -23.6201
14.5648 | -34.7167 | -23.6282
14.5709 | -34.7171 | -23.6363
14.577 | -34.7176 | -23.6445
14.5831 | -34.718 | -23.6528
14.5892 | -34.7185 | -23.6611
14.5953 | -34.719 | -23.6692
14.6014 | -34.7195 | -23.6775
14.6075 | -34.7199 | -23.6856
14.6136 | -34.7204 | -23.6939
14.6197 | -34.7209 | -23.7023
14.6258 | -34.7214 | -23.7106
14.6319 | -34.7218 | -23.7189
14.638 | -34.7223 | -23.7271
14.6441 | -34.7227 | -23.7355
14.6502 | -34.7232 | -23.7438
14.6563 | -34.7237 | -23.7521
14.6624 | -34.7242 | -23.7604
14.6685 | -34.7247 | -23.7684
14.6746 | -34.7252 | -23.7768
14.6807 | -34.7256 | -23.785
14.6869 | -34.7261 | -23.7933
14.693 | -34.7266 | -23.8015
14.6991 | -34.7271 | -23.8097
14.7052 | -34.7276 | -23.8178
14.7113 | -34.728 | -23.8261
14.7174 | -34.7285 | -23.8344
14.7235 | -34.729 | -23.8425
14.7296 | -34.7295 | -23.8508
14.7357 | -34.73 | -23.859
14.7418 | -34.7305 | -23.8671
14.7479 | -34.7309 | -23.8753
14.754 | -34.7314 | -23.8836
14.7601 | -34.7319 | -23.8916
14.7662 | -34.7324 | -23.8998
14.7723 | -34.7329 | -23.9077
14.7784 | -34.7334 | -23.916
14.7845 | -34.7338 | -23.9242
14.7906 | -34.7343 | -23.9326
14.7967 | -34.7348 | -23.9407
14.8028 | -34.7352 | -23.9489
14.8089 | -34.7357 | -23.957
14.815 | -34.7362 | -23.9653
14.8211 | -34.7367 | -23.9735
14.8272 | -34.7372 | -23.9817
14.8334 | -34.7377 | -23.9899
14.8395 | -34.7381 | -23.998
14.8456 | -34.7386 | -24.0061
14.8517 | -34.7391 | -24.0145
14.8578 | -34.7396 | -24.0227
14.8639 | -34.74 | -24.0309
14.87 | -34.7405 | -24.0391
14.8761 | -34.741 | -24.0473
14.8822 | -34.7415 | -24.0555
14.8883 | -34.742 | -24.0638
14.8944 | -34.7424 | -24.0718
14.9005 | -34.7429 | -24.08
14.9066 | -34.7434 | -24.0882
14.9127 | -34.7439 | -24.0963
14.9188 | -34.7443 | -24.1045
14.9249 | -34.7448 | -24.1127
14.931 | -34.7453 | -24.121
14.9371 | -34.7458 | -24.1291
14.9432 | -34.7463 | -24.1373
14.9493 | -34.7467 | -24.1455
14.9554 | -34.7472 | -24.1536
14.9615 | -34.7477 | -24.1619
14.9676 | -34.7481 | -24.1702
14.9738 | -34.7486 | -24.1783
14.9799 | -34.7491 | -24.1865
14.986 | -34.7496 | -24.1945
14.9921 | -34.7501 | -24.2026
14.9982 | -34.7505 | -24.2108
15.0043 | -34.751 | -24.2189
15.0104 | -34.7515 | -24.2273
15.0165 | -34.752 | -24.2355
15.0226 | -34.7525 | -24.2436
15.0287 | -34.753 | -24.2518
15.0348 | -34.7535 | -24.26
15.0409 | -34.754 | -24.2681
15.047 | -34.7544 | -24.2764
15.0531 | -34.7549 | -24.2845
15.0592 | -34.7554 | -24.2926
15.0653 | -34.7559 | -24.3006
15.0714 | -34.7563 | -24.3087
15.0775 | -34.7568 | -24.3167
15.0836 | -34.7573 | -24.3249
15.0897 | -34.7578 | -24.3329
15.0958 | -34.7583 | -24.341
15.1019 | -34.7588 | -24.3492
15.108 | -34.7593 | -24.3573
15.1141 | -34.7598 | -24.3654
15.1203 | -34.7602 | -24.3736
15.1264 | -34.7607 | -24.3817
15.1325 | -34.7612 | -24.3899
15.1386 | -34.7617 | -24.3981
15.1447 | -34.7622 | -24.4062
15.1508 | -34.7627 | -24.4144
15.1569 | -34.7632 | -24.4225
15.163 | -34.7637 | -24.4305
15.1691 | -34.7642 | -24.4386
15.1752 | -34.7646 | -24.4468
15.1813 | -34.7651 | -24.455
15.1874 | -34.7656 | -24.4632
15.1935 | -34.7661 | -24.4712
15.1996 | -34.7666 | -24.4793
15.2057 | -34.7671 | -24.4874
15.2118 | -34.7675 | -24.4956
15.2179 | -34.768 | -24.5037
15.224 | -34.7685 | -24.5118
15.2301 | -34.769 | -24.5198
15.2362 | -34.7695 | -24.5279
15.2423 | -34.77 | -24.5361
15.2484 | -34.7705 | -24.5443
15.2545 | -34.771 | -24.5525
15.2607 | -34.7715 | -24.5606
15.2668 | -34.772 | -24.5685
15.2729 | -34.7725 | -24.5765
15.279 | -34.773 | -24.585
15.2851 | -34.7735 | -24.5929
15.2912 | -34.774 | -24.601
15.2973 | -34.7745 | -24.6092
15.3034 | -34.775 | -24.6173
15.3095 | -34.7755 | -24.6254
15.3156 | -34.776 | -24.6334
15.3217 | -34.7765 | -24.6415
15.3278 | -34.777 | -24.6498
15.3339 | -34.7774 | -24.6576
15.34 | -34.7779 | -24.6658
15.3461 | -34.7784 | -24.6738
15.3522 | -34.7789 | -24.682
15.3583 | -34.7794 | -24.6902
15.3644 | -34.7799 | -24.6981
15.3705 | -34.7804 | -24.7058
15.3766 | -34.7809 | -24.7141
15.3827 | -34.7814 | -24.7221
15.3888 | -34.7819 | -24.7302
15.3949 | -34.7824 | -24.7384
15.401 | -34.7828 | -24.7465
15.4072 | -34.7833 | -24.7544
15.4133 | -34.7838 | -24.7625
15.4194 | -34.7843 | -24.7706
15.4255 | -34.7848 | -24.7786
15.4316 | -34.7853 | -24.7867
15.4377 | -34.7858 | -24.7947
15.4438 | -34.7863 | -24.8027
15.4499 | -34.7868 | -24.8107
15.456 | -34.7873 | -24.8187
15.4621 | -34.7878 | -24.8266
15.4682 | -34.7883 | -24.8347
15.4743 | -34.7888 | -24.8426
15.4804 | -34.7893 | -24.8507
15.4865 | -34.7897 | -24.8588
15.4926 | -34.7902 | -24.8669
15.4987 | -34.7907 | -24.8749
15.5048 | -34.7912 | -24.883
15.5109 | -34.7917 | -24.891
15.517 | -34.7922 | -24.899
15.5231 | -34.7927 | -24.9072
15.5292 | -34.7932 | -24.9152
15.5353 | -34.7937 | -24.9234
15.5414 | -34.7942 | -24.9314
15.5476 | -34.7947 | -24.9392
15.5537 | -34.7952 | -24.9471
15.5598 | -34.7957 | -24.9552
15.5659 | -34.7962 | -24.9632
15.572 | -34.7967 | -24.9712
15.5781 | -34.7971 | -24.9792
15.5842 | -34.7976 | -24.9871
15.5903 | -34.7981 | -24.9951
15.5964 | -34.7987 | -25.0031
15.6025 | -34.7991 | -25.011
15.6086 | -34.7996 | -25.0191
15.6147 | -34.8001 | -25.0273
15.6208 | -34.8007 | -25.035
15.6269 | -34.8012 | -25.0431
15.633 | -34.8017 | -25.0511
15.6391 | -34.8022 | -25.0591
15.6452 | -34.8027 | -25.0672
15.6513 | -34.8032 | -25.0749
15.6574 | -34.8037 | -25.083
15.6635 | -34.8042 | -25.091
15.6696 | -34.8047 | -25.099
15.6757 | -34.8051 | -25.1072
15.6818 | -34.8056 | -25.1152
15.688 | -34.8061 | -25.1229
15.6941 | -34.8066 | -25.1308
15.7002 | -34.8071 | -25.1387
15.7063 | -34.8076 | -25.1467
15.7124 | -34.8081 | -25.1547
15.7185 | -34.8086 | -25.1626
15.7246 | -34.8091 | -25.1704
15.7307 | -34.8096 | -25.1783
15.7368 | -34.8101 | -25.1862
15.7429 | -34.8106 | -25.1941
15.749 | -34.8111 | -25.2021
15.7551 | -34.8115 | -25.21
15.7612 | -34.8121 | -25.2178
15.7673 | -34.8126 | -25.2257
15.7734 | -34.8131 | -25.2336
15.7795 | -34.8135 | -25.2416
15.7856 | -34.814 | -25.2493
15.7917 | -34.8145 | -25.2571
15.7978 | -34.815 | -25.2652
15.8039 | -34.8155 | -25.2733
15.81 | -34.8161 | -25.2812
15.8161 | -34.8165 | -25.2891
15.8222 | -34.8171 | -25.2969
15.8283 | -34.8176 | -25.3047
15.8345 | -34.8181 | -25.3127
15.8406 | -34.8186 | -25.3206
15.8467 | -34.8191 | -25.3286
15.8528 | -34.8196 | -25.3365
15.8589 | -34.8201 | -25.3445
15.865 | -34.8206 | -25.3522
15.8711 | -34.8211 | -25.3601
15.8772 | -34.8216 | -25.3678
15.8833 | -34.8221 | -25.3758
15.8894 | -34.8226 | -25.3838
15.8955 | -34.823 | -25.3917
15.9016 | -34.8236 | -25.3996
15.9077 | -34.824 | -25.4074
15.9138 | -34.8245 | -25.4154
15.9199 | -34.825 | -25.4233
15.926 | -34.8255 | -25.431
15.9321 | -34.826 | -25.4387
15.9382 | -34.8265 | -25.4464
15.9443 | -34.827 | -25.4543
15.9504 | -34.8275 | -25.4622
15.9565 | -34.828 | -25.4702
15.9626 | -34.8284 | -25.4782
15.9687 | -34.8289 | -25.4858
15.9749 | -34.8294 | -25.4936
15.981 | -34.8299 | -25.5015
15.9871 | -34.8304 | -25.5095
15.9932 | -34.8309 | -25.5172
15.9993 | -34.8314 | -25.5249
16.0054 | -34.8319 | -25.5325
16.0115 | -34.8324 | -25.5404
16.0176 | -34.8329 | -25.5483
16.0237 | -34.8334 | -25.5561
16.0298 | -34.8338 | -25.5639
16.0359 | -34.8343 | -25.5716
16.042 | -34.8348 | -25.5794
16.0481 | -34.8353 | -25.5872
16.0542 | -34.8358 | -25.595
16.0603 | -34.8363 | -25.6026
16.0664 | -34.8368 | -25.6106
16.0725 | -34.8373 | -25.6181
16.0786 | -34.8378 | -25.6258
16.0847 | -34.8383 | -25.6337
16.0908 | -34.8388 | -25.6415
16.0969 | -34.8393 | -25.6494
16.103 | -34.8398 | -25.6572
16.1091 | -34.8403 | -25.6648
16.1152 | -34.8408 | -25.6726
16.1214 | -34.8413 | -25.6804
16.1275 | -34.8418 | -25.6882
16.1336 | -34.8423 | -25.6959
16.1397 | -34.8428 | -25.7037
16.1458 | -34.8433 | -25.7114
16.1519 | -34.8438 | -25.7192
16.158 | -34.8443 | -25.7272
16.1641 | -34.8448 | -25.735
16.1702 | -34.8453 | -25.7428
16.1763 | -34.8458 | -25.7506
16.1824 | -34.8463 | -25.758
16.1885 | -34.8468 | -25.7656
16.1946 | -34.8473 | -25.7732
16.2007 | -34.8478 | -25.781
16.2068 | -34.8483 | -25.7886
16.2129 | -34.8488 | -25.7965
16.219 | -34.8493 | -25.8039
16.2251 | -34.8498 | -25.8116
16.2312 | -34.8503 | -25.8193
16.2373 | -34.8507 | -25.8271
16.2434 | -34.8513 | -25.835
16.2495 | -34.8517 | -25.8426
16.2556 | -34.8522 | -25.8505
16.2618 | -34.8527 | -25.8582
16.2679 | -34.8533 | -25.866
16.274 | -34.8538 | -25.8738
16.2801 | -34.8543 | -25.8817
16.2862 | -34.8548 | -25.8891
16.2923 | -34.8553 | -25.8968
16.2984 | -34.8557 | -25.9045
16.3045 | -34.8562 | -25.9123
16.3106 | -34.8567 | -25.9201
16.3167 | -34.8572 | -25.928
16.3228 | -34.8577 | -25.9357
16.3289 | -34.8582 | -25.9433
16.335 | -34.8587 | -25.9511
16.3411 | -34.8591 | -25.9589
16.3472 | -34.8597 | -25.9668
16.3533 | -34.8601 | -25.9747
16.3594 | -34.8606 | -25.9822
16.3655 | -34.8611 | -25.9898
16.3716 | -34.8616 | -25.9975
16.3777 | -34.8621 | -26.005
16.3838 | -34.8626 | -26.0128
16.3899 | -34.8631 | -26.0201
16.396 | -34.8636 | -26.0278
16.4021 | -34.8641 | -26.0354
16.4083 | -34.8646 | -26.043
16.4144 | -34.8651 | -26.0509
16.4205 | -34.8656 | -26.0585
16.4266 | -34.8661 | -26.0662
16.4327 | -34.8666 | -26.0737
16.4388 | -34.8671 | -26.0815
16.4449 | -34.8676 | -26.089
16.451 | -34.8681 | -26.0967
16.4571 | -34.8685 | -26.1044
16.4632 | -34.8691 | -26.1118
16.4693 | -34.8695 | -26.1195
16.4754 | -34.87 | -26.1272
16.4815 | -34.8705 | -26.1349
16.4876 | -34.871 | -26.1424
16.4937 | -34.8715 | -26.1498
16.4998 | -34.872 | -26.1574
16.5059 | -34.8725 | -26.165
16.512 | -34.873 | -26.1728
16.5181 | -34.8735 | -26.1806
16.5242 | -34.874 | -26.1884
16.5303 | -34.8745 | -26.196
16.5364 | -34.875 | -26.2037
16.5425 | -34.8755 | -26.2113
16.5487 | -34.876 | -26.219
16.5548 | -34.8764 | -26.2266
16.5609 | -34.877 | -26.2343
16.567 | -34.8775 | -26.2417
16.5731 | -34.878 | -26.2494
16.5792 | -34.8785 | -26.257
16.5853 | -34.879 | -26.2647
16.5914 | -34.8795 | -26.2724
16.5975 | -34.8799 | -26.2802
16.6036 | -34.8804 | -26.2878
16.6097 | -34.8809 | -26.2954
16.6158 | -34.8814 | -26.303
16.6219 | -34.8819 | -26.3105
16.628 | -34.8824 | -26.3182
16.6341 | -34.8829 | -26.3258
16.6402 | -34.8834 | -26.3332
16.6463 | -34.8838 | -26.3409
16.6524 | -34.8844 | -26.3486
16.6585 | -34.8849 | -26.3563
16.6646 | -34.8854 | -26.3637
16.6707 | -34.8858 | -26.3714
16.6768 | -34.8864 | -26.3788
16.6829 | -34.8868 | -26.3862
16.689 | -34.8874 | -26.3939
16.6952 | -34.8879 | -26.4015
16.7013 | -34.8883 | -26.4092
16.7074 | -34.8888 | -26.4166
16.7135 | -34.8893 | -26.4242
16.7196 | -34.8898 | -26.4318
16.7257 | -34.8903 | -26.4394
16.7318 | -34.8908 | -26.4469
16.7379 | -34.8913 | -26.4545
16.744 | -34.8918 | -26.4618
16.7501 | -34.8923 | -26.4694
16.7562 | -34.8928 | -26.477
16.7623 | -34.8932 | -26.4845
16.7684 | -34.8937 | -26.492
16.7745 | -34.8942 | -26.4998
16.7806 | -34.8947 | -26.5072
16.7867 | -34.8952 | -26.5147
16.7928 | -34.8957 | -26.5223
16.7989 | -34.8962 | -26.5298
16.805 | -34.8967 | -26.5374
16.8111 | -34.8972 | -26.5449
16.8172 | -34.8977 | -26.5525
16.8233 | -34.8982 | -26.5601
16.8294 | -34.8987 | -26.5676
16.8356 | -34.8992 | -26.5752
16.8417 | -34.8997 | -26.5827
16.8478 | -34.9002 | -26.5901
16.8539 | -34.9007 | -26.5977
16.86 | -34.9011 | -26.6052
16.8661 | -34.9017 | -26.6128
16.8722 | -34.9021 | -26.6206
16.8783 | -34.9026 | -26.6282
16.8844 | -34.9031 | -26.6356
16.8905 | -34.9036 | -26.6433
16.8966 | -34.9041 | -26.6507
16.9027 | -34.9046 | -26.6583
16.9088 | -34.905 | -26.6659
16.9149 | -34.9055 | -26.6734
16.921 | -34.906 | -26.6809
16.9271 | -34.9065 | -26.6883
16.9332 | -34.907 | -26.6958
16.9393 | -34.9075 | -26.7035
16.9454 | -34.908 | -26.7108
16.9515 | -34.9085 | -26.7185
16.9576 | -34.909 | -26.7256
16.9637 | -34.9095 | -26.733
16.9698 | -34.91 | -26.7408
16.9759 | -34.9105 | -26.7481
16.9821 | -34.911 | -26.7556
16.9882 | -34.9115 | -26.7628
16.9943 | -34.912 | -26.7703
17.0004 | -34.9125 | -26.7776
17.0065 | -34.913 | -26.7852
17.0126 | -34.9135 | -26.7928
17.0187 | -34.9139 | -26.8003
17.0248 | -34.9144 | -26.8079
17.0309 | -34.9149 | -26.8156
17.037 | -34.9154 | -26.823
17.0431 | -34.9159 | -26.8306
17.0492 | -34.9164 | -26.8379
17.0553 | -34.9168 | -26.8452
17.0614 | -34.9173 | -26.8529
17.0675 | -34.9178 | -26.8603
17.0736 | -34.9183 | -26.868
17.0797 | -34.9188 | -26.8752
17.0858 | -34.9193 | -26.8829
17.0919 | -34.9198 | -26.8906
17.098 | -34.9203 | -26.8978
17.1041 | -34.9208 | -26.9054
17.1102 | -34.9213 | -26.9129
17.1163 | -34.9218 | -26.9202
17.1225 | -34.9223 | -26.9276
17.1286 | -34.9228 | -26.9349
17.1347 | -34.9233 | -26.9425
17.1408 | -34.9238 | -26.9502
17.1469 | -34.9243 | -26.9577
17.153 | -34.9248 | -26.9651
17.1591 | -34.9253 | -26.9725
17.1652 | -34.9258 | -26.9797
17.1713 | -34.9263 | -26.9871
17.1774 | -34.9268 | -26.9946
17.1835 | -34.9273 | -27.0019
17.1896 | -34.9277 | -27.0093
17.1957 | -34.9282 | -27.0167
17.2018 | -34.9287 | -27.0238
17.2079 | -34.9292 | -27.0313
17.214 | -34.9297 | -27.0389
17.2201 | -34.9302 | -27.0463
17.2262 | -34.9307 | -27.0539
17.2323 | -34.9312 | -27.0614
17.2384 | -34.9317 | -27.0687
17.2445 | -34.9322 | -27.076
17.2506 | -34.9327 | -27.0834
17.2567 | -34.9332 | -27.0909
17.2628 | -34.9337 | -27.0985
17.269 | -34.9342 | -27.1058
17.2751 | -34.9347 | -27.1134
17.2812 | -34.9352 | -27.1209
17.2873 | -34.9357 | -27.1284
17.2934 | -34.9362 | -27.136
17.2995 | -34.9367 | -27.1435
17.3056 | -34.9372 | -27.1508
17.3117 | -34.9376 | -27.1583
17.3178 | -34.9381 | -27.1656
17.3239 | -34.9386 | -27.1729
17.33 | -34.9391 | -27.1802
17.3361 | -34.9396 | -27.1876
17.3422 | -34.94 | -27.1949
17.3483 | -34.9405 | -27.2024
17.3544 | -34.941 | -27.21
17.3605 | -34.9415 | -27.2174
17.3666 | -34.942 | -27.2249
17.3727 | -34.9425 | -27.2323
17.3788 | -34.943 | -27.2394
17.3849 | -34.9435 | -27.2467
17.391 | -34.944 | -27.2542
17.3971 | -34.9444 | -27.2614
17.4032 | -34.945 | -27.2686
17.4094 | -34.9454 | -27.2758
17.4155 | -34.9459 | -27.2831
17.4216 | -34.9464 | -27.2904
17.4277 | -34.9469 | -27.298
17.4338 | -34.9474 | -27.3054
17.4399 | -34.9479 | -27.3131
17.446 | -34.9484 | -27.3203
17.4521 | -34.9489 | -27.3276
17.4582 | -34.9494 | -27.3348
17.4643 | -34.9499 | -27.3422
17.4704 | -34.9504 | -27.3495
17.4765 | -34.9509 | -27.3569
17.4826 | -34.9514 | -27.3641
17.4887 | -34.9518 | -27.3715
17.4948 | -34.9523 | -27.3788
17.5009 | -34.9528 | -27.3862
17.507 | -34.9533 | -27.3936
17.5131 | -34.9538 | -27.4009
17.5192 | -34.9543 | -27.408
17.5253 | -34.9548 | -27.4155
17.5314 | -34.9552 | -27.4228
17.5375 | -34.9557 | -27.4304
17.5436 | -34.9562 | -27.4378
17.5497 | -34.9567 | -27.4453
17.5559 | -34.9571 | -27.4522
17.562 | -34.9577 | -27.4596
17.5681 | -34.9582 | -27.467
17.5742 | -34.9587 | -27.4745
17.5803 | -34.9592 | -27.4818
17.5864 | -34.9597 | -27.489
17.5925 | -34.9602 | -27.4963
17.5986 | -34.9607 | -27.5035
17.6047 | -34.9611 | -27.511
17.6108 | -34.9617 | -27.5184
17.6169 | -34.9621 | -27.5256
17.623 | -34.9627 | -27.5328
17.6291 | -34.9631 | -27.5401
17.6352 | -34.9636 | -27.5473
17.6413 | -34.9641 | -27.5548
17.6474 | -34.9646 | -27.5621
17.6535 | -34.9651 | -27.5696
17.6596 | -34.9655 | -27.5771
17.6657 | -34.966 | -27.5846
17.6718 | -34.9665 | -27.5921
17.6779 | -34.967 | -27.5994
17.684 | -34.9675 | -27.6068
17.6901 | -34.968 | -27.6142
17.6963 | -34.9684 | -27.6213
17.7024 | -34.969 | -27.6284
17.7085 | -34.9694 | -27.6356
17.7146 | -34.9699 | -27.6428
17.7207 | -34.9704 | -27.6499
17.7268 | -34.9709 | -27.6571
17.7329 | -34.9714 | -27.6645
17.739 | -34.9719 | -27.6719
17.7451 | -34.9724 | -27.679
17.7512 | -34.9729 | -27.6862
17.7573 | -34.9734 | -27.6935
17.7634 | -34.9738 | -27.7007
17.7695 | -34.9743 | -27.7081
17.7756 | -34.9748 | -27.7154
17.7817 | -34.9753 | -27.7227
17.7878 | -34.9758 | -27.7302
17.7939 | -34.9763 | -27.7375
17.8 | -34.9767 | -27.7449
17.8061 | -34.9773 | -27.7522
17.8122 | -34.9778 | -27.7595
17.8183 | -34.9783 | -27.7667
17.8244 | -34.9788 | -27.7741
17.8305 | -34.9793 | -27.7815
17.8366 | -34.9798 | -27.7887
17.8428 | -34.9802 | -27.796
17.8489 | -34.9807 | -27.8033
17.855 | -34.9812 | -27.8103
17.8611 | -34.9817 | -27.8175
17.8672 | -34.9822 | -27.8247
17.8733 | -34.9827 | -27.8321
17.8794 | -34.9832 | -27.8393
17.8855 | -34.9837 | -27.8466
17.8916 | -34.9842 | -27.8539
17.8977 | -34.9846 | -27.8612
17.9038 | -34.9851 | -27.8682
17.9099 | -34.9856 | -27.8755
17.916 | -34.9861 | -27.8828
17.9221 | -34.9866 | -27.8901
17.9282 | -34.9871 | -27.8972
17.9343 | -34.9876 | -27.9046
17.9404 | -34.988 | -27.9117
17.9465 | -34.9885 | -27.9189
17.9526 | -34.989 | -27.9261
17.9587 | -34.9895 | -27.9334
17.9648 | -34.99 | -27.9405
17.9709 | -34.9904 | -27.9478
17.977 | -34.9909 | -27.9552
17.9832 | -34.9914 | -27.9626
17.9893 | -34.9919 | -27.9699
17.9954 | -34.9924 | -27.9773
18.0015 | -34.9929 | -27.9845
18.0076 | -34.9934 | -27.9918
18.0137 | -34.9938 | -27.999
18.0198 | -34.9943 | -28.0062
18.0259 | -34.9948 | -28.0133
18.032 | -34.9953 | -28.0205
18.0381 | -34.9958 | -28.0279
18.0442 | -34.9963 | -28.0351
18.0503 | -34.9968 | -28.0425
18.0564 | -34.9973 | -28.0498
18.0625 | -34.9977 | -28.057
18.0686 | -34.9982 | -28.0643
18.0747 | -34.9987 | -28.0715
18.0808 | -34.9992 | -28.0785
18.0869 | -34.9997 | -28.0858
18.093 | -35.0002 | -28.0931
18.0991 | -35.0007 | -28.1004
18.1052 | -35.0011 | -28.1078
18.1113 | -35.0016 | -28.1148
18.1174 | -35.0021 | -28.122
18.1236 | -35.0026 | -28.1293
18.1297 | -35.0031 | -28.1366
18.1358 | -35.0036 | -28.1438
18.1419 | -35.0041 | -28.151
18.148 | -35.0046 | -28.1584
18.1541 | -35.0051 | -28.1656
18.1602 | -35.0056 | -28.1729
18.1663 | -35.0061 | -28.1801
18.1724 | -35.0066 | -28.1873
18.1785 | -35.0071 | -28.1944
18.1846 | -35.0075 | -28.2018
18.1907 | -35.008 | -28.209
18.1968 | -35.0085 | -28.2161
18.2029 | -35.009 | -28.2233
18.209 | -35.0095 | -28.2307
18.2151 | -35.01 | -28.238
18.2212 | -35.0105 | -28.245
18.2273 | -35.011 | -28.2523
18.2334 | -35.0114 | -28.2595
18.2395 | -35.0119 | -28.2667
18.2456 | -35.0124 | -28.274
18.2517 | -35.0129 | -28.2813
18.2578 | -35.0133 | -28.2884
18.2639 | -35.0139 | -28.2958
18.2701 | -35.0143 | -28.3031
18.2762 | -35.0149 | -28.3103
18.2823 | -35.0153 | -28.3175
18.2884 | -35.0158 | -28.3247
18.2945 | -35.0163 | -28.3318
18.3006 | -35.0168 | -28.3392
18.3067 | -35.0173 | -28.3464
18.3128 | -35.0178 | -28.3538
18.3189 | -35.0183 | -28.361
18.325 | -35.0187 | -28.3681
18.3311 | -35.0192 | -28.3754
18.3372 | -35.0197 | -28.3826
18.3433 | -35.0202 | -28.39
18.3494 | -35.0207 | -28.3973
18.3555 | -35.0212 | -28.4046
18.3616 | -35.0218 | -28.4117
18.3677 | -35.0223 | -28.4191
18.3738 | -35.0227 | -28.4262
18.3799 | -35.0232 | -28.4333
18.386 | -35.0237 | -28.4406
18.3921 | -35.0241 | -28.4478
18.3982 | -35.0246 | -28.4548
18.4043 | -35.0251 | -28.462
18.4105 | -35.0256 | -28.4693
18.4166 | -35.0261 | -28.4766
18.4227 | -35.0265 | -28.4838
18.4288 | -35.027 | -28.4911
18.4349 | -35.0276 | -28.4984
18.441 | -35.028 | -28.5055
18.4471 | -35.0285 | -28.5126
18.4532 | -35.029 | -28.52
18.4593 | -35.0295 | -28.5271
18.4654 | -35.0299 | -28.5342
18.4715 | -35.0304 | -28.5415
18.4776 | -35.0309 | -28.5488
18.4837 | -35.0314 | -28.556
18.4898 | -35.0319 | -28.5634
18.4959 | -35.0324 | -28.5707
18.502 | -35.0329 | -28.5779
18.5081 | -35.0333 | -28.5852
18.5142 | -35.0338 | -28.5925
18.5203 | -35.0343 | -28.5996
18.5264 | -35.0347 | -28.6069
18.5325 | -35.0352 | -28.6142
18.5386 | -35.0357 | -28.6215
18.5447 | -35.0362 | -28.6286
18.5508 | -35.0367 | -28.6358
18.557 | -35.0372 | -28.6431
18.5631 | -35.0376 | -28.6503
18.5692 | -35.0381 | -28.6574
18.5753 | -35.0386 | -28.6647
18.5814 | -35.0391 | -28.6718
18.5875 | -35.0396 | -28.679
18.5936 | -35.04 | -28.6862
18.5997 | -35.0405 | -28.6935
18.6058 | -35.041 | -28.7005
18.6119 | -35.0415 | -28.7077
18.618 | -35.042 | -28.7149
18.6241 | -35.0424 | -28.7221
18.6302 | -35.0429 | -28.7294
18.6363 | -35.0434 | -28.7368
18.6424 | -35.0439 | -28.7439
18.6485 | -35.0443 | -28.751
18.6546 | -35.0448 | -28.7582
18.6607 | -35.0453 | -28.7653
18.6668 | -35.0458 | -28.7727
18.6729 | -35.0463 | -28.7801
18.679 | -35.0467 | -28.7871
18.6851 | -35.0472 | -28.7941
18.6912 | -35.0477 | -28.8013
18.6974 | -35.0482 | -28.8084
18.7035 | -35.0487 | -28.8156
18.7096 | -35.0491 | -28.8228
18.7157 | -35.0496 | -28.83
18.7218 | -35.0501 | -28.8373
18.7279 | -35.0506 | -28.8446
18.734 | -35.0511 | -28.852
18.7401 | -35.0516 | -28.8591
18.7462 | -35.0521 | -28.8662
18.7523 | -35.0526 | -28.8735
18.7584 | -35.0531 | -28.8806
18.7645 | -35.0535 | -28.8879
18.7706 | -35.054 | -28.8949
18.7767 | -35.0545 | -28.9022
18.7828 | -35.055 | -28.9093
18.7889 | -35.0555 | -28.9164
18.795 | -35.0559 | -28.9236
18.8011 | -35.0564 | -28.9308
18.8072 | -35.0569 | -28.9381
18.8133 | -35.0573 | -28.9453
18.8194 | -35.0578 | -28.9521
18.8255 | -35.0583 | -28.9593
18.8316 | -35.0588 | -28.9665
18.8377 | -35.0593 | -28.9735
18.8439 | -35.0597 | -28.9806
18.85 | -35.0602 | -28.9877
18.8561 | -35.0607 | -28.995
18.8622 | -35.0612 | -29.0022
18.8683 | -35.0616 | -29.0094
18.8744 | -35.0621 | -29.0166
18.8805 | -35.0626 | -29.0239
18.8866 | -35.0631 | -29.0312
18.8927 | -35.0636 | -29.0382
18.8988 | -35.0641 | -29.0455
18.9049 | -35.0646 | -29.0525
18.911 | -35.065 | -29.0599
18.9171 | -35.0655 | -29.0669
18.9232 | -35.066 | -29.0738
18.9293 | -35.0665 | -29.0808
18.9354 | -35.067 | -29.0878
18.9415 | -35.0675 | -29.095
18.9476 | -35.068 | -29.102
18.9537 | -35.0685 | -29.1092
18.9598 | -35.069 | -29.1162
18.9659 | -35.0694 | -29.1233
18.972 | -35.0699 | -29.1304
18.9781 | -35.0704 | -29.1375
18.9843 | -35.0708 | -29.1446
18.9904 | -35.0713 | -29.1517
18.9965 | -35.0718 | -29.1586
19.0026 | -35.0723 | -29.1659
19.0087 | -35.0728 | -29.1732
19.0148 | -35.0733 | -29.1802
19.0209 | -35.0738 | -29.1875
19.027 | -35.0743 | -29.1947
19.0331 | -35.0748 | -29.2015
19.0392 | -35.0752 | -29.2086
19.0453 | -35.0757 | -29.2157
19.0514 | -35.0762 | -29.2227
19.0575 | -35.0766 | -29.2299
19.0636 | -35.0771 | -29.2367
19.0697 | -35.0776 | -29.2439
19.0758 | -35.0781 | -29.251
19.0819 | -35.0785 | -29.258
19.088 | -35.079 | -29.2651
19.0941 | -35.0795 | -29.2724
19.1002 | -35.08 | -29.2795
19.1063 | -35.0805 | -29.2866
19.1124 | -35.0809 | -29.2937
19.1185 | -35.0814 | -29.3009
19.1246 | -35.0818 | -29.308
19.1308 | -35.0823 | -29.3153
19.1369 | -35.0828 | -29.3224
19.143 | -35.0833 | -29.3296
19.1491 | -35.0837 | -29.337
19.1552 | -35.0842 | -29.3442
19.1613 | -35.0846 | -29.3514
19.1674 | -35.0851 | -29.3585
19.1735 | -35.0856 | -29.3657
19.1796 | -35.0861 | -29.3726
19.1857 | -35.0865 | -29.3797
19.1918 | -35.087 | -29.3869
19.1979 | -35.0875 | -29.3941
19.204 | -35.088 | -29.4013
19.2101 | -35.0885 | -29.4083
19.2162 | -35.089 | -29.4155
19.2223 | -35.0895 | -29.4227
19.2284 | -35.0899 | -29.4299
19.2345 | -35.0904 | -29.437
19.2406 | -35.0909 | -29.4438
19.2467 | -35.0914 | -29.4509
19.2528 | -35.0918 | -29.4582
19.2589 | -35.0923 | -29.4653
19.265 | -35.0928 | -29.4724
19.2712 | -35.0932 | -29.4794
19.2773 | -35.0937 | -29.4865
19.2834 | -35.0942 | -29.4935
19.2895 | -35.0947 | -29.5007
19.2956 | -35.0952 | -29.5079
19.3017 | -35.0957 | -29.5149
19.3078 | -35.0961 | -29.5221
19.3139 | -35.0966 | -29.529
19.32 | -35.0971 | -29.5361
19.3261 | -35.0976 | -29.5433
19.3322 | -35.098 | -29.5502
19.3383 | -35.0985 | -29.5573
19.3444 | -35.0989 | -29.5643
19.3505 | -35.0994 | -29.5714
19.3566 | -35.0999 | -29.5785
19.3627 | -35.1004 | -29.5854
19.3688 | -35.1009 | -29.5925
19.3749 | -35.1014 | -29.5996
19.381 | -35.1019 | -29.6064
19.3871 | -35.1024 | -29.6133
19.3932 | -35.1028 | -29.6204
19.3993 | -35.1033 | -29.6274
19.4054 | -35.1038 | -29.6345
19.4115 | -35.1043 | -29.6418
19.4177 | -35.1047 | -29.6488
19.4238 | -35.1052 | -29.6557
19.4299 | -35.1057 | -29.6627
19.436 | -35.1061 | -29.6698
19.4421 | -35.1066 | -29.6768
19.4482 | -35.1071 | -29.6838
19.4543 | -35.1076 | -29.6909
19.4604 | -35.108 | -29.698
19.4665 | -35.1085 | -29.7048
19.4726 | -35.109 | -29.7117
19.4787 | -35.1094 | -29.7186
19.4848 | -35.1099 | -29.7257
19.4909 | -35.1103 | -29.7324
19.497 | -35.1108 | -29.7393
19.5031 | -35.1113 | -29.7463
19.5092 | -35.1117 | -29.7535
19.5153 | -35.1122 | -29.7604
19.5214 | -35.1126 | -29.7673
19.5275 | -35.1131 | -29.7743
19.5336 | -35.1136 | -29.7813
19.5397 | -35.114 | -29.7883
19.5458 | -35.1145 | -29.7955
19.5519 | -35.1149 | -29.8025
19.5581 | -35.1154 | -29.8092
19.5642 | -35.1159 | -29.8163
19.5703 | -35.1164 | -29.8233
19.5764 | -35.1168 | -29.8303
19.5825 | -35.1173 | -29.8374
19.5886 | -35.1178 | -29.8445
19.5947 | -35.1183 | -29.8512
19.6008 | -35.1187 | -29.8583
19.6069 | -35.1192 | -29.8654
19.613 | -35.1197 | -29.8726
19.6191 | -35.1201 | -29.8797
19.6252 | -35.1206 | -29.8866
19.6313 | -35.121 | -29.8936
19.6374 | -35.1215 | -29.9006
19.6435 | -35.1219 | -29.9077
19.6496 | -35.1224 | -29.9147
19.6557 | -35.1229 | -29.9218
19.6618 | -35.1233 | -29.9287
19.6679 | -35.1238 | -29.9358
19.674 | -35.1243 | -29.9428
19.6801 | -35.1247 | -29.9498
19.6862 | -35.1252 | -29.9569
19.6923 | -35.1257 | -29.9639
19.6984 | -35.1262 | -29.9707
19.7046 | -35.1267 | -29.9778
19.7107 | -35.1272 | -29.9848
19.7168 | -35.1276 | -29.9917
19.7229 | -35.1281 | -29.9988
19.729 | -35.1285 | -30.0059
19.7351 | -35.129 | -30.0126
19.7412 | -35.1294 | -30.0199
19.7473 | -35.1299 | -30.0267
19.7534 | -35.1303 | -30.0339
19.7595 | -35.1308 | -30.0412
19.7656 | -35.1312 | -30.0483
19.7717 | -35.1317 | -30.055
19.7778 | -35.1321 | -30.0618
19.7839 | -35.1326 | -30.069
19.79 | -35.133 | -30.0761
19.7961 | -35.1335 | -30.0833
19.8022 | -35.1339 | -30.0901
19.8083 | -35.1344 | -30.097
19.8144 | -35.1349 | -30.1038
19.8205 | -35.1354 | -30.1109
19.8266 | -35.1359 | -30.1177
19.8327 | -35.1363 | -30.1247
19.8388 | -35.1368 | -30.1312
19.845 | -35.1372 | -30.1382
19.8511 | -35.1377 | -30.1451
19.8572 | -35.1382 | -30.1521
19.8633 | -35.1386 | -30.1592
19.8694 | -35.1391 | -30.1664
19.8755 | -35.1395 | -30.1727
19.8816 | -35.14 | -30.1798
19.8877 | -35.1404 | -30.1867
19.8938 | -35.1409 | -30.1934
19.8999 | -35.1413 | -30.2002
19.906 | -35.1418 | -30.2072
19.9121 | -35.1423 | -30.2139
19.9182 | -35.1427 | -30.2209
19.9243 | -35.1432 | -30.2279
19.9304 | -35.1437 | -30.2349
19.9365 | -35.1441 | -30.2417
19.9426 | -35.1446 | -30.2486
19.9487 | -35.1451 | -30.2557
19.9548 | -35.1455 | -30.2625
19.9609 | -35.1461 | -30.2694
19.967 | -35.1465 | -30.2764
19.9731 | -35.147 | -30.2833
19.9792 | -35.1474 | -30.2898
19.9853 | -35.1479 | -30.2966
19.9915 | -35.1484 | -30.3035
19.9976 | -35.1488 | -30.3105
20.0037 | -35.1493 | -30.3172
20.0098 | -35.1498 | -30.3242
20.0159 | -35.1503 | -30.331
20.022 | -35.1507 | -30.3379
20.0281 | -35.1512 | -30.3448
20.0342 | -35.1517 | -30.3519
20.0403 | -35.1522 | -30.3586
20.0464 | -35.1526 | -30.3654
20.0525 | -35.1531 | -30.3721
20.0586 | -35.1536 | -30.3791
20.0647 | -35.1541 | -30.3859
20.0708 | -35.1545 | -30.393
20.0769 | -35.155 | -30.4002
20.083 | -35.1554 | -30.4069
20.0891 | -35.1559 | -30.414
20.0952 | -35.1564 | -30.4212
20.1013 | -35.1568 | -30.428
20.1074 | -35.1573 | -30.4349
20.1135 | -35.1578 | -30.4418
20.1196 | -35.1582 | -30.4482
20.1257 | -35.1586 | -30.455
20.1319 | -35.1591 | -30.4618
20.138 | -35.1595 | -30.4683
20.1441 | -35.16 | -30.4751
20.1502 | -35.1605 | -30.4821
20.1563 | -35.1609 | -30.4885
20.1624 | -35.1613 | -30.4955
20.1685 | -35.1617 | -30.5022
20.1746 | -35.1622 | -30.5091
20.1807 | -35.1626 | -30.516
20.1868 | -35.1631 | -30.5228
20.1929 | -35.1635 | -30.5297
20.199 | -35.164 | -30.5367
20.2051 | -35.1645 | -30.5434
20.2112 | -35.1649 | -30.5503
20.2173 | -35.1654 | -30.5571
20.2234 | -35.1658 | -30.564
20.2295 | -35.1663 | -30.5709
20.2356 | -35.1668 | -30.5777
20.2417 | -35.1672 | -30.5845
20.2478 | -35.1677 | -30.5912
20.2539 | -35.1681 | -30.5979
20.26 | -35.1686 | -30.6044
20.2661 | -35.1691 | -30.6115
20.2723 | -35.1696 | -30.6185
20.2784 | -35.17 | -30.6252
20.2845 | -35.1705 | -30.6319
20.2906 | -35.171 | -30.6388
20.2967 | -35.1714 | -30.6455
20.3028 | -35.1719 | -30.6522
20.3089 | -35.1724 | -30.659
20.315 | -35.1728 | -30.6661
20.3211 | -35.1733 | -30.673
20.3272 | -35.1737 | -30.68
20.3333 | -35.1742 | -30.6864
20.3394 | -35.1746 | -30.6932
20.3455 | -35.1751 | -30.6998
20.3516 | -35.1755 | -30.7064
20.3577 | -35.176 | -30.7131
20.3638 | -35.1764 | -30.7198
20.3699 | -35.1769 | -30.7261
20.376 | -35.1773 | -30.733
20.3821 | -35.1778 | -30.7399
20.3882 | -35.1782 | -30.7466
20.3943 | -35.1787 | -30.7534
20.4004 | -35.1791 | -30.7596
20.4065 | -35.1795 | -30.7663
20.4126 | -35.18 | -30.773
20.4188 | -35.1805 | -30.7799
20.4249 | -35.1809 | -30.7866
20.431 | -35.1814 | -30.7934
20.4371 | -35.1818 | -30.7998
20.4432 | -35.1822 | -30.8066
20.4493 | -35.1827 | -30.8134
20.4554 | -35.1831 | -30.8201
20.4615 | -35.1836 | -30.8269
20.4676 | -35.184 | -30.8338
20.4737 | -35.1845 | -30.8404
20.4798 | -35.185 | -30.8472
20.4859 | -35.1855 | -30.8541
20.492 | -35.1859 | -30.8606
20.4981 | -35.1864 | -30.8673
20.5042 | -35.1868 | -30.8741
20.5103 | -35.1872 | -30.8803
20.5164 | -35.1876 | -30.8872
20.5225 | -35.1881 | -30.8943
20.5286 | -35.1885 | -30.9011
20.5347 | -35.189 | -30.9077
20.5408 | -35.1895 | -30.9141
20.5469 | -35.1899 | -30.9207
20.553 | -35.1904 | -30.9274
20.5592 | -35.1908 | -30.9341
20.5653 | -35.1913 | -30.9408
20.5714 | -35.1918 | -30.9476
20.5775 | -35.1922 | -30.9539
20.5836 | -35.1927 | -30.9608
20.5897 | -35.1931 | -30.9675
20.5958 | -35.1935 | -30.9742
20.6019 | -35.194 | -30.9812
20.608 | -35.1944 | -30.9879
20.6141 | -35.1949 | -30.9946
20.6202 | -35.1953 | -31.0011
20.6263 | -35.1958 | -31.0078
20.6324 | -35.1962 | -31.0145
20.6385 | -35.1967 | -31.0212
20.6446 | -35.1971 | -31.0278
20.6507 | -35.1975 | -31.0347
20.6568 | -35.198 | -31.0415
20.6629 | -35.1984 | -31.0483
20.669 | -35.1989 | -31.055
20.6751 | -35.1993 | -31.0615
20.6812 | -35.1998 | -31.0679
20.6873 | -35.2002 | -31.0747
20.6934 | -35.2006 | -31.0815
20.6995 | -35.2011 | -31.0881
20.7057 | -35.2015 | -31.0949
20.7118 | -35.202 | -31.1017
20.7179 | -35.2024 | -31.1083
20.724 | -35.2029 | -31.115
20.7301 | -35.2033 | -31.1217
20.7362 | -35.2038 | -31.1286
20.7423 | -35.2042 | -31.1353
20.7484 | -35.2047 | -31.1419
20.7545 | -35.2052 | -31.1486
20.7606 | -35.2056 | -31.1553
20.7667 | -35.2061 | -31.1619
20.7728 | -35.2066 | -31.1683
20.7789 | -35.207 | -31.1748
20.785 | -35.2074 | -31.1818
20.7911 | -35.2078 | -31.1881
20.7972 | -35.2082 | -31.1948
20.8033 | -35.2087 | -31.2018
20.8094 | -35.2092 | -31.2085
20.8155 | -35.2096 | -31.215
20.8216 | -35.21 | -31.2215
20.8277 | -35.2104 | -31.2277
20.8338 | -35.2109 | -31.2341
20.8399 | -35.2114 | -31.2409
20.8461 | -35.2119 | -31.2476
20.8522 | -35.2123 | -31.254
20.8583 | -35.2128 | -31.2602
20.8644 | -35.2132 | -31.2672
20.8705 | -35.2136 | -31.2739
20.8766 | -35.2141 | -31.2802
20.8827 | -35.2145 | -31.287
20.8888 | -35.215 | -31.2935
20.8949 | -35.2154 | -31.3002
20.901 | -35.2158 | -31.307
20.9071 | -35.2163 | -31.3132
20.9132 | -35.2168 | -31.3197
20.9193 | -35.2173 | -31.3265
20.9254 | -35.2177 | -31.333
20.9315 | -35.2182 | -31.3397
20.9376 | -35.2186 | -31.3462
20.9437 | -35.2191 | -31.3529
20.9498 | -35.2195 | -31.3596
20.9559 | -35.2199 | -31.3662
20.962 | -35.2204 | -31.3728
20.9681 | -35.2208 | -31.3789
20.9742 | -35.2213 | -31.3857
20.9803 | -35.2218 | -31.3922
20.9864 | -35.2222 | -31.3988
20.9926 | -35.2227 | -31.4055
20.9987 | -35.2231 | -31.4121
21.0048 | -35.2236 | -31.4186
21.0109 | -35.224 | -31.4252
21.017 | -35.2245 | -31.4319
21.0231 | -35.225 | -31.4384
21.0292 | -35.2254 | -31.445
21.0353 | -35.2258 | -31.4512
21.0414 | -35.2263 | -31.4577
21.0475 | -35.2267 | -31.4644
21.0536 | -35.2271 | -31.4706
21.0597 | -35.2275 | -31.4772
21.0658 | -35.2279 | -31.4838
21.0719 | -35.2284 | -31.4902
21.078 | -35.2288 | -31.497
21.0841 | -35.2292 | -31.5037
21.0902 | -35.2297 | -31.5099
21.0963 | -35.2299 | -31.516
21.1024 | -35.2304 | -31.5227
21.1085 | -35.2308 | -31.529
21.1146 | -35.2313 | -31.5352
21.1207 | -35.2317 | -31.5416
21.1268 | -35.2321 | -31.5482
21.133 | -35.2326 | -31.555
21.1391 | -35.2331 | -31.5615
21.1452 | -35.2334 | -31.568
21.1513 | -35.2338 | -31.5744
21.1574 | -35.2342 | -31.5808
21.1635 | -35.2347 | -31.5875
21.1696 | -35.2351 | -31.5938
21.1757 | -35.2356 | -31.5998
21.1818 | -35.236 | -31.6063
21.1879 | -35.2364 | -31.6129
21.194 | -35.2369 | -31.6197
21.2001 | -35.2373 | -31.6262
21.2062 | -35.2378 | -31.6326
21.2123 | -35.2382 | -31.639
21.2184 | -35.2387 | -31.6456
21.2245 | -35.2392 | -31.6519
21.2306 | -35.2396 | -31.6585
21.2367 | -35.2401 | -31.6651
21.2428 | -35.2405 | -31.6711
21.2489 | -35.241 | -31.6776
21.255 | -35.2414 | -31.6842
21.2611 | -35.2418 | -31.6907
21.2672 | -35.2422 | -31.697
21.2733 | -35.2427 | -31.7036
21.2795 | -35.2431 | -31.7095
21.2856 | -35.2435 | -31.7159
21.2917 | -35.2439 | -31.7225
21.2978 | -35.2443 | -31.729
21.3039 | -35.2448 | -31.7359
21.31 | -35.2453 | -31.7426
21.3161 | -35.2457 | -31.7486
21.3222 | -35.2461 | -31.755
21.3283 | -35.2466 | -31.7613
21.3344 | -35.247 | -31.7677
21.3405 | -35.2474 | -31.7738
21.3466 | -35.2479 | -31.7802
21.3527 | -35.2483 | -31.7866
21.3588 | -35.2487 | -31.7931
21.3649 | -35.2492 | -31.7992
21.371 | -35.2496 | -31.8059
21.3771 | -35.25 | -31.8128
21.3832 | -35.2505 | -31.8193
21.3893 | -35.2509 | -31.8259
21.3954 | -35.2513 | -31.8323
21.4015 | -35.2517 | -31.8385
21.4076 | -35.2521 | -31.845
21.4137 | -35.2525 | -31.8511
21.4199 | -35.253 | -31.8573
21.426 | -35.2534 | -31.8631
21.4321 | -35.2538 | -31.8696
21.4382 | -35.2542 | -31.8762
21.4443 | -35.2547 | -31.8825
21.4504 | -35.2551 | -31.8889
21.4565 | -35.2556 | -31.8945
21.4626 | -35.256 | -31.901
21.4687 | -35.2564 | -31.9076
21.4748 | -35.2568 | -31.914
21.4809 | -35.2573 | -31.9205
21.487 | -35.2577 | -31.9268
21.4931 | -35.2581 | -31.933
21.4992 | -35.2585 | -31.9395
21.5053 | -35.259 | -31.9459
21.5114 | -35.2594 | -31.9524
21.5175 | -35.2598 | -31.9585
21.5236 | -35.2602 | -31.965
21.5297 | -35.2606 | -31.971
21.5358 | -35.261 | -31.9771
21.5419 | -35.2615 | -31.9836
21.548 | -35.2619 | -31.9899
21.5541 | -35.2623 | -31.9965
21.5602 | -35.2627 | -32.0024
21.5664 | -35.2631 | -32.008
21.5725 | -35.2636 | -32.0143
21.5786 | -35.2641 | -32.0205
21.5847 | -35.2645 | -32.0269
21.5908 | -35.2649 | -32.0331
21.5969 | -35.2653 | -32.0387
21.603 | -35.2658 | -32.0453
21.6091 | -35.2662 | -32.0517
21.6152 | -35.2666 | -32.0581
21.6213 | -35.2671 | -32.0643
21.6274 | -35.2675 | -32.0706
21.6335 | -35.268 | -32.0767
21.6396 | -35.2684 | -32.0832
21.6457 | -35.2689 | -32.0896
21.6518 | -35.2693 | -32.096
21.6579 | -35.2697 | -32.1027
21.664 | -35.2701 | -32.1093
21.6701 | -35.2706 | -32.1156
21.6762 | -35.271 | -32.1216
21.6823 | -35.2714 | -32.128
21.6884 | -35.2719 | -32.1344
21.6945 | -35.2723 | -32.1411
21.7006 | -35.2728 | -32.1476
21.7068 | -35.2732 | -32.1537
21.7129 | -35.2736 | -32.1598
21.719 | -35.2741 | -32.1664
21.7251 | -35.2745 | -32.1726
21.7312 | -35.2749 | -32.1791
21.7373 | -35.2753 | -32.185
21.7434 | -35.2757 | -32.1915
21.7495 | -35.2761 | -32.1979
21.7556 | -35.2765 | -32.2045
21.7617 | -35.277 | -32.2112
21.7678 | -35.2774 | -32.2173
21.7739 | -35.2779 | -32.2233
21.78 | -35.2784 | -32.2297
21.7861 | -35.2788 | -32.2361
21.7922 | -35.2792 | -32.2425
21.7983 | -35.2797 | -32.2486
21.8044 | -35.2801 | -32.2553
21.8105 | -35.2806 | -32.2614
21.8166 | -35.281 | -32.2681
21.8227 | -35.2814 | -32.2739
21.8288 | -35.2819 | -32.2801
21.8349 | -35.2823 | -32.2864
21.841 | -35.2827 | -32.2928
21.8471 | -35.2831 | -32.2989
21.8533 | -35.2835 | -32.3055
21.8594 | -35.284 | -32.3122
21.8655 | -35.2844 | -32.3185
21.8716 | -35.2849 | -32.3249
21.8777 | -35.2853 | -32.3307
21.8838 | -35.2858 | -32.337
21.8899 | -35.2862 | -32.3433
21.896 | -35.2866 | -32.3494
21.9021 | -35.2871 | -32.3555
21.9082 | -35.2875 | -32.3619
21.9143 | -35.2879 | -32.3679
21.9204 | -35.2884 | -32.3737
21.9265 | -35.2888 | -32.3799
21.9326 | -35.2892 | -32.3862
21.9387 | -35.2897 | -32.3927
21.9448 | -35.2901 | -32.3992
21.9509 | -35.2905 | -32.4056
21.957 | -35.291 | -32.4119
21.9631 | -35.2915 | -32.4182
21.9692 | -35.2919 | -32.4244
21.9753 | -35.2923 | -32.431
21.9814 | -35.2927 | -32.4375
21.9875 | -35.2931 | -32.4435
21.9937 | -35.2935 | -32.4497
21.9998 | -35.2939 | -32.4556
22.0059 | -35.2943 | -32.4618
22.012 | -35.2948 | -32.4682
22.0181 | -35.2952 | -32.4743
22.0242 | -35.2957 | -32.4804
22.0303 | -35.2961 | -32.4866
22.0364 | -35.2965 | -32.4929
22.0425 | -35.2969 | -32.499
22.0486 | -35.2973 | -32.5053
22.0547 | -35.2977 | -32.5115
22.0608 | -35.2981 | -32.5179
22.0669 | -35.2985 | -32.5242
22.073 | -35.299 | -32.5305
22.0791 | -35.2995 | -32.5365
22.0852 | -35.2999 | -32.5428
22.0913 | -35.3003 | -32.5489
22.0974 | -35.3007 | -32.5551
22.1035 | -35.3011 | -32.5611
22.1096 | -35.3016 | -32.5676
22.1157 | -35.302 | -32.574
22.1218 | -35.3024 | -32.5801
22.1279 | -35.3028 | -32.5861
22.134 | -35.3033 | -32.5924
22.1402 | -35.3037 | -32.5983
22.1463 | -35.3041 | -32.6048
22.1524 | -35.3045 | -32.611
22.1585 | -35.305 | -32.6172
22.1646 | -35.3054 | -32.6231
22.1707 | -35.3059 | -32.6293
22.1768 | -35.3063 | -32.6358
22.1829 | -35.3067 | -32.6417
22.189 | -35.3071 | -32.648
22.1951 | -35.3075 | -32.6537
22.2012 | -35.3079 | -32.6597
22.2073 | -35.3084 | -32.6657
22.2134 | -35.3088 | -32.6715
22.2195 | -35.3092 | -32.6777
22.2256 | -35.3096 | -32.6838
22.2317 | -35.3101 | -32.6897
22.2378 | -35.3105 | -32.6957
22.2439 | -35.311 | -32.7018
22.25 | -35.3114 | -32.7077
22.2561 | -35.3118 | -32.7139
22.2622 | -35.3122 | -32.72
22.2683 | -35.3126 | -32.7257
22.2744 | -35.313 | -32.7316
22.2806 | -35.3134 | -32.7377
22.2867 | -35.3138 | -32.7439
22.2928 | -35.3142 | -32.7504
22.2989 | -35.3146 | -32.7563
22.305 | -35.3151 | -32.7624
22.3111 | -35.3155 | -32.7686
22.3172 | -35.3159 | -32.7748
22.3233 | -35.3163 | -32.7813
22.3294 | -35.3168 | -32.7877
22.3355 | -35.3173 | -32.7942
22.3416 | -35.3177 | -32.8006
22.3477 | -35.3181 | -32.8068
22.3538 | -35.3185 | -32.8129
22.3599 | -35.3189 | -32.8187
22.366 | -35.3193 | -32.8251
22.3721 | -35.3197 | -32.8312
22.3782 | -35.3201 | -32.8377
22.3843 | -35.3206 | -32.8432
22.3904 | -35.3209 | -32.849
22.3965 | -35.3213 | -32.8556
22.4026 | -35.3216 | -32.8621
22.4087 | -35.322 | -32.8674
22.4148 | -35.3225 | -32.874
22.4209 | -35.3229 | -32.8806
22.4271 | -35.3233 | -32.8863
22.4332 | -35.3238 | -32.8926
22.4393 | -35.3241 | -32.8985
22.4454 | -35.3245 | -32.9036
22.4515 | -35.325 | -32.9092
22.4576 | -35.3254 | -32.9151
22.4637 | -35.3257 | -32.9214
22.4698 | -35.3262 | -32.9278
22.4759 | -35.3265 | -32.934
22.482 | -35.3268 | -32.9399
22.4881 | -35.3273 | -32.9458
22.4942 | -35.3277 | -32.9521
22.5003 | -35.3281 | -32.9581
22.5064 | -35.3285 | -32.9645
22.5125 | -35.3289 | -32.9708
22.5186 | -35.3293 | -32.9772
22.5247 | -35.3297 | -32.9836
22.5308 | -35.3301 | -32.99
22.5369 | -35.3305 | -32.9959
22.543 | -35.3309 | -33.0023
22.5491 | -35.3314 | -33.0077
22.5552 | -35.3317 | -33.0135
22.5613 | -35.3322 | -33.0197
22.5675 | -35.3326 | -33.026
22.5736 | -35.333 | -33.032
22.5797 | -35.3335 | -33.0382
22.5858 | -35.334 | -33.0437
22.5919 | -35.3344 | -33.0493
22.598 | -35.3348 | -33.0556
22.6041 | -35.3352 | -33.0615
22.6102 | -35.3356 | -33.0676
22.6163 | -35.336 | -33.0735
22.6224 | -35.3363 | -33.0797
22.6285 | -35.3367 | -33.0864
22.6346 | -35.3371 | -33.0929
22.6407 | -35.3375 | -33.099
22.6468 | -35.3379 | -33.1053
22.6529 | -35.3383 | -33.1116
22.659 | -35.3387 | -33.1178
22.6651 | -35.3392 | -33.1239
22.6712 | -35.3396 | -33.1294
22.6773 | -35.34 | -33.1353
22.6834 | -35.3405 | -33.1409
22.6895 | -35.341 | -33.1461
22.6956 | -35.3415 | -33.1526
22.7017 | -35.3419 | -33.1584
22.7079 | -35.3423 | -33.1643
22.714 | -35.3427 | -33.1701
22.7201 | -35.3431 | -33.176
22.7262 | -35.3434 | -33.1818
22.7323 | -35.3438 | -33.1879
22.7384 | -35.3443 | -33.1939
22.7445 | -35.3447 | -33.2001
22.7506 | -35.3451 | -33.2062
22.7567 | -35.3455 | -33.2126
22.7628 | -35.3458 | -33.2181
22.7689 | -35.3463 | -33.224
22.775 | -35.3466 | -33.2303
22.7811 | -35.347 | -33.2365
22.7872 | -35.3474 | -33.2429
22.7933 | -35.3478 | -33.2488
22.7994 | -35.3482 | -33.2548
22.8055 | -35.3486 | -33.2612
22.8116 | -35.349 | -33.2671
22.8177 | -35.3494 | -33.2735
22.8238 | -35.3499 | -33.2794
22.8299 | -35.3503 | -33.285
22.836 | -35.3507 | -33.2912
22.8421 | -35.3512 | -33.2974
22.8482 | -35.3516 | -33.3039
22.8544 | -35.352 | -33.3098
22.8605 | -35.3524 | -33.3159
22.8666 | -35.3528 | -33.3217
22.8727 | -35.3532 | -33.3279
22.8788 | -35.3536 | -33.3339
22.8849 | -35.354 | -33.3398
22.891 | -35.3544 | -33.3457
22.8971 | -35.3548 | -33.3514
22.9032 | -35.3551 | -33.3572
22.9093 | -35.3555 | -33.363
22.9154 | -35.3559 | -33.369
22.9215 | -35.3563 | -33.3748
22.9276 | -35.3567 | -33.3806
22.9337 | -35.3571 | -33.3861
22.9398 | -35.3575 | -33.3919
22.9459 | -35.358 | -33.3982
22.952 | -35.3583 | -33.4044
22.9581 | -35.3587 | -33.4099
22.9642 | -35.3591 | -33.4159
22.9703 | -35.3596 | -33.4219
22.9764 | -35.36 | -33.4278
22.9825 | -35.3604 | -33.4338
22.9886 | -35.3608 | -33.4398
22.9948 | -35.3613 | -33.4459
23.0009 | -35.3617 | -33.4522
23.007 | -35.3622 | -33.458
23.0131 | -35.3626 | -33.4638
23.0192 | -35.363 | -33.4698
23.0253 | -35.3634 | -33.4762
23.0314 | -35.3637 | -33.4823
23.0375 | -35.3642 | -33.4882
23.0436 | -35.3647 | -33.4937
23.0497 | -35.3651 | -33.5
23.0558 | -35.3655 | -33.5057
23.0619 | -35.3659 | -33.5119
23.068 | -35.3664 | -33.5175
23.0741 | -35.3668 | -33.5223
23.0802 | -35.3672 | -33.5281
23.0863 | -35.3676 | -33.5342
23.0924 | -35.3681 | -33.5399
23.0985 | -35.3684 | -33.546
23.1046 | -35.3688 | -33.552
23.1107 | -35.3691 | -33.5579
23.1168 | -35.3695 | -33.5633
23.1229 | -35.3699 | -33.5695
23.129 | -35.3704 | -33.5754
23.1351 | -35.3708 | -33.5818
23.1413 | -35.3712 | -33.5876
23.1474 | -35.3717 | -33.5934
23.1535 | -35.3721 | -33.5995
23.1596 | -35.3724 | -33.6056
23.1657 | -35.3728 | -33.6116
23.1718 | -35.3732 | -33.6176
23.1779 | -35.3736 | -33.6232
23.184 | -35.374 | -33.6293
23.1901 | -35.3744 | -33.6351
23.1962 | -35.3748 | -33.6411
23.2023 | -35.3752 | -33.6472
23.2084 | -35.3756 | -33.6535
23.2145 | -35.376 | -33.659
23.2206 | -35.3764 | -33.6649
23.2267 | -35.3768 | -33.6711
23.2328 | -35.3772 | -33.6769
23.2389 | -35.3776 | -33.6827
23.245 | -35.3781 | -33.6886
23.2511 | -35.3785 | -33.6943
23.2572 | -35.379 | -33.7003
23.2633 | -35.3794 | -33.706
23.2694 | -35.3798 | -33.712
23.2755 | -35.3802 | -33.7182
23.2817 | -35.3806 | -33.7242
23.2878 | -35.3809 | -33.7295
23.2939 | -35.3813 | -33.7358
23.3 | -35.3817 | -33.7416
23.3061 | -35.382 | -33.7477
23.3122 | -35.3825 | -33.7537
23.3183 | -35.3829 | -33.7595
23.3244 | -35.3833 | -33.7651
23.3305 | -35.3837 | -33.7711
23.3366 | -35.3841 | -33.7769
23.3427 | -35.3845 | -33.783
23.3488 | -35.3849 | -33.7889
23.3549 | -35.3853 | -33.7949
23.361 | -35.3857 | -33.8003
23.3671 | -35.3861 | -33.806
23.3732 | -35.3866 | -33.8122
23.3793 | -35.387 | -33.8183
23.3854 | -35.3874 | -33.8245
23.3915 | -35.3878 | -33.8304
23.3976 | -35.3882 | -33.8365
23.4037 | -35.3886 | -33.8422
23.4098 | -35.389 | -33.8482
23.4159 | -35.3894 | -33.8543
23.422 | -35.3899 | -33.8605
23.4282 | -35.3903 | -33.8663
23.4343 | -35.3907 | -33.8722
23.4404 | -35.3911 | -33.8779
23.4465 | -35.3915 | -33.8841
23.4526 | -35.3918 | -33.8899
23.4587 | -35.3922 | -33.8955
23.4648 | -35.3927 | -33.9005
23.4709 | -35.393 | -33.9061
23.477 | -35.3933 | -33.9118
23.4831 | -35.3938 | -33.9174
23.4892 | -35.3942 | -33.9231
23.4953 | -35.3946 | -33.9289
23.5014 | -35.3949 | -33.9345
23.5075 | -35.3953 | -33.9405
23.5136 | -35.3957 | -33.9461
23.5197 | -35.3962 | -33.952
23.5258 | -35.3965 | -33.9577
23.5319 | -35.397 | -33.9631
23.538 | -35.3973 | -33.9689
23.5441 | -35.3977 | -33.9749
23.5502 | -35.3981 | -33.9807
23.5563 | -35.3985 | -33.9868
23.5624 | -35.3989 | -33.9923
23.5686 | -35.3993 | -33.9986
23.5747 | -35.3997 | -34.0045
23.5808 | -35.4001 | -34.0102
23.5869 | -35.4005 | -34.0159
23.593 | -35.401 | -34.0216
23.5991 | -35.4014 | -34.028
23.6052 | -35.4017 | -34.0344
23.6113 | -35.4021 | -34.0402
23.6174 | -35.4025 | -34.0461
23.6235 | -35.4029 | -34.0518
23.6296 | -35.4033 | -34.0578
23.6357 | -35.4037 | -34.0641
23.6418 | -35.4041 | -34.0701
23.6479 | -35.4045 | -34.0757
23.654 | -35.405 | -34.0812
23.6601 | -35.4054 | -34.0868
23.6662 | -35.4059 | -34.0924
23.6723 | -35.4063 | -34.0977
23.6784 | -35.4067 | -34.1039
23.6845 | -35.4071 | -34.1096
23.6906 | -35.4075 | -34.1154
23.6967 | -35.4078 | -34.1212
23.7028 | -35.4082 | -34.1271
23.7089 | -35.4086 | -34.1331
23.7151 | -35.409 | -34.1386
23.7212 | -35.4093 | -34.1444
23.7273 | -35.4097 | -34.1501
23.7334 | -35.4101 | -34.156
23.7395 | -35.4105 | -34.1618
23.7456 | -35.4109 | -34.1675
23.7517 | -35.4113 | -34.1732
23.7578 | -35.4117 | -34.179
23.7639 | -35.4121 | -34.1848
23.77 | -35.4125 | -34.191
23.7761 | -35.4128 | -34.1969
23.7822 | -35.4131 | -34.2026
23.7883 | -35.4135 | -34.2086
23.7944 | -35.4139 | -34.2142
23.8005 | -35.4143 | -34.2205
23.8066 | -35.4147 | -34.2266
23.8127 | -35.4151 | -34.2323
23.8188 | -35.4155 | -34.2381
23.8249 | -35.4159 | -34.2439
23.831 | -35.4163 | -34.2497
23.8371 | -35.4167 | -34.2556
23.8432 | -35.4171 | -34.2609
23.8493 | -35.4175 | -34.2669
23.8555 | -35.4179 | -34.2728
23.8616 | -35.4183 | -34.2785
23.8677 | -35.4187 | -34.2844
23.8738 | -35.419 | -34.2908
23.8799 | -35.4194 | -34.2969
23.886 | -35.4198 | -34.3023
23.8921 | -35.4201 | -34.3083
23.8982 | -35.4205 | -34.3145
23.9043 | -35.421 | -34.3203
23.9104 | -35.4213 | -34.3258
23.9165 | -35.4218 | -34.3309
23.9226 | -35.4222 | -34.3374
23.9287 | -35.4226 | -34.3437
23.9348 | -35.423 | -34.3499
23.9409 | -35.4234 | -34.3561
23.947 | -35.4238 | -34.3613
23.9531 | -35.424 | -34.3672
23.9592 | -35.4244 | -34.3734
23.9653 | -35.4248 | -34.3789
23.9714 | -35.4252 | -34.3845
23.9775 | -35.4256 | -34.3904
23.9836 | -35.426 | -34.3967
23.9897 | -35.4264 | -34.403
23.9958 | -35.4269 | -34.4089
24.002 | -35.4273 | -34.4148
24.0081 | -35.4277 | -34.4211
24.0142 | -35.428 | -34.4264
24.0203 | -35.4285 | -34.432
24.0264 | -35.4288 | -34.4377
24.0325 | -35.4293 | -34.4436
24.0386 | -35.4296 | -34.4495
24.0447 | -35.4301 | -34.455
24.0508 | -35.4305 | -34.4603
24.0569 | -35.4308 | -34.4662
24.063 | -35.4313 | -34.4714
24.0691 | -35.4317 | -34.477
24.0752 | -35.432 | -34.4822
24.0813 | -35.4324 | -34.4878
24.0874 | -35.4327 | -34.4941
24.0935 | -35.4331 | -34.5005
24.0996 | -35.4334 | -34.5059
24.1057 | -35.4338 | -34.5118
24.1118 | -35.4342 | -34.5171
24.1179 | -35.4347 | -34.5233
24.124 | -35.4351 | -34.5286
24.1301 | -35.4354 | -34.5336
24.1362 | -35.4358 | -34.5395
24.1424 | -35.4361 | -34.5458
24.1485 | -35.4365 | -34.5516
24.1546 | -35.4369 | -34.5575
24.1607 | -35.4372 | -34.5628
24.1668 | -35.4376 | -34.5682
24.1729 | -35.438 | -34.5736
24.179 | -35.4384 | -34.5795
24.1851 | -35.4387 | -34.5853
24.1912 | -35.439 | -34.591
24.1973 | -35.4395 | -34.5967
24.2034 | -35.4399 | -34.6025
24.2095 | -35.4402 | -34.608
24.2156 | -35.4406 | -34.6135
24.2217 | -35.441 | -34.6193
24.2278 | -35.4414 | -34.6248
24.2339 | -35.4418 | -34.6302
24.24 | -35.4421 | -34.6359
24.2461 | -35.4425 | -34.6414
24.2522 | -35.4429 | -34.6475
24.2583 | -35.4433 | -34.6533
24.2644 | -35.4438 | -34.6591
24.2705 | -35.4442 | -34.6642
24.2766 | -35.4446 | -34.6699
24.2827 | -35.445 | -34.6758
24.2889 | -35.4454 | -34.6817
24.295 | -35.4457 | -34.6874
24.3011 | -35.4462 | -34.693
24.3072 | -35.4467 | -34.6987
24.3133 | -35.447 | -34.7043
24.3194 | -35.4474 | -34.7104
24.3255 | -35.4478 | -34.7162
24.3316 | -35.4482 | -34.722
24.3377 | -35.4486 | -34.7276
24.3438 | -35.4488 | -34.7327
24.3499 | -35.4492 | -34.7384
24.356 | -35.4496 | -34.7439
24.3621 | -35.45 | -34.7498
24.3682 | -35.4503 | -34.7556
24.3743 | -35.4507 | -34.7616
24.3804 | -35.4512 | -34.7674
24.3865 | -35.4515 | -34.7729
24.3926 | -35.4519 | -34.7784
24.3987 | -35.4522 | -34.7838
24.4048 | -35.4526 | -34.7889
24.4109 | -35.453 | -34.7942
24.417 | -35.4534 | -34.8004
24.4231 | -35.4538 | -34.8055
24.4293 | -35.4542 | -34.8113
24.4354 | -35.4546 | -34.8168
24.4415 | -35.455 | -34.8226
24.4476 | -35.4554 | -34.8282
24.4537 | -35.4558 | -34.8336
24.4598 | -35.4562 | -34.8392
24.4659 | -35.4566 | -34.8446
24.472 | -35.457 | -34.8502
24.4781 | -35.4574 | -34.8559
24.4842 | -35.4577 | -34.8615
24.4903 | -35.4581 | -34.8675
24.4964 | -35.4585 | -34.8731
24.5025 | -35.4588 | -34.8786
24.5086 | -35.4592 | -34.8839
24.5147 | -35.4596 | -34.8894
24.5208 | -35.46 | -34.8945
24.5269 | -35.4603 | -34.9002
24.533 | -35.4608 | -34.9057
24.5391 | -35.4612 | -34.9112
24.5452 | -35.4616 | -34.9172
24.5513 | -35.462 | -34.9227
24.5574 | -35.4624 | -34.9283
24.5635 | -35.4628 | -34.9339
24.5696 | -35.4631 | -34.9396
24.5758 | -35.4635 | -34.9452
24.5819 | -35.4639 | -34.9506
24.588 | -35.4643 | -34.9564
24.5941 | -35.4647 | -34.9621
24.6002 | -35.4651 | -34.9679
24.6063 | -35.4655 | -34.9734
24.6124 | -35.4659 | -34.9788
24.6185 | -35.4663 | -34.9845
24.6246 | -35.4667 | -34.9901
24.6307 | -35.4671 | -34.996
24.6368 | -35.4675 | -35.0018
24.6429 | -35.4678 | -35.0073
24.649 | -35.4683 | -35.0128
24.6551 | -35.4686 | -35.0183
24.6612 | -35.469 | -35.0234
24.6673 | -35.4694 | -35.0291
24.6734 | -35.4698 | -35.0344
24.6795 | -35.4702 | -35.0401
24.6856 | -35.4706 | -35.0457
24.6917 | -35.471 | -35.0516
24.6978 | -35.4714 | -35.0569
24.7039 | -35.4718 | -35.0625
24.71 | -35.4722 | -35.068
24.7162 | -35.4725 | -35.0737
24.7223 | -35.4729 | -35.0796
24.7284 | -35.4733 | -35.0847
24.7345 | -35.4736 | -35.0902
24.7406 | -35.474 | -35.0959
24.7467 | -35.4744 | -35.1017
24.7528 | -35.4747 | -35.1075
24.7589 | -35.4751 | -35.1128
24.765 | -35.4754 | -35.1178
24.7711 | -35.4758 | -35.1235
24.7772 | -35.4761 | -35.1292
24.7833 | -35.4764 | -35.1349
24.7894 | -35.4769 | -35.1408
24.7955 | -35.4772 | -35.1468
24.8016 | -35.4775 | -35.1524
24.8077 | -35.4779 | -35.1586
24.8138 | -35.4783 | -35.1644
24.8199 | -35.4787 | -35.17
24.826 | -35.4791 | -35.176
24.8321 | -35.4795 | -35.1819
24.8382 | -35.4798 | -35.187
24.8443 | -35.4801 | -35.193
24.8504 | -35.4805 | -35.1987
24.8565 | -35.4808 | -35.2047
24.8627 | -35.4813 | -35.2104
24.8688 | -35.4817 | -35.2166
24.8749 | -35.482 | -35.2226
24.881 | -35.4823 | -35.2282
24.8871 | -35.4828 | -35.2341
24.8932 | -35.4832 | -35.2397
24.8993 | -35.4835 | -35.2456
24.9054 | -35.4839 | -35.2513
24.9115 | -35.4842 | -35.2567
24.9176 | -35.4846 | -35.2622
24.9237 | -35.485 | -35.2676
24.9298 | -35.4854 | -35.2735
24.9359 | -35.4859 | -35.2794
24.942 | -35.4863 | -35.2847
24.9481 | -35.4867 | -35.2903
24.9542 | -35.487 | -35.2958
24.9603 | -35.4873 | -35.3014
24.9664 | -35.4876 | -35.307
24.9725 | -35.488 | -35.3126
24.9786 | -35.4884 | -35.3182
24.9847 | -35.4887 | -35.3233
24.9908 | -35.4891 | -35.3289
24.9969 | -35.4894 | -35.3346
25.0031 | -35.4899 | -35.3406
25.0092 | -35.4903 | -35.3465
25.0153 | -35.4907 | -35.3519
25.0214 | -35.491 | -35.3577
25.0275 | -35.4914 | -35.3631
25.0336 | -35.4917 | -35.3687
25.0397 | -35.492 | -35.3742
25.0458 | -35.4924 | -35.3803
25.0519 | -35.4928 | -35.3866
25.058 | -35.4932 | -35.3921
25.0641 | -35.4935 | -35.3973
25.0702 | -35.4938 | -35.4022
25.0763 | -35.4942 | -35.4074
25.0824 | -35.4944 | -35.4126
25.0885 | -35.4948 | -35.4179
25.0946 | -35.4952 | -35.4237
25.1007 | -35.4956 | -35.4292
25.1068 | -35.496 | -35.4348
25.1129 | -35.4964 | -35.4403
25.119 | -35.4968 | -35.4462
25.1251 | -35.4973 | -35.4516
25.1312 | -35.4977 | -35.4573
25.1373 | -35.4981 | -35.4631
25.1435 | -35.4985 | -35.469
25.1496 | -35.4988 | -35.4745
25.1557 | -35.4992 | -35.48
25.1618 | -35.4996 | -35.4857
25.1679 | -35.5 | -35.4912
25.174 | -35.5003 | -35.4972
25.1801 | -35.5007 | -35.5023
25.1862 | -35.501 | -35.5074
25.1923 | -35.5014 | -35.513
25.1984 | -35.5019 | -35.519
25.2045 | -35.5022 | -35.5247
25.2106 | -35.5026 | -35.5299
25.2167 | -35.5029 | -35.5359
25.2228 | -35.5032 | -35.5411
25.2289 | -35.5036 | -35.5462
25.235 | -35.5039 | -35.5518
25.2411 | -35.5043 | -35.5576
25.2472 | -35.5047 | -35.5631
25.2533 | -35.505 | -35.5688
25.2594 | -35.5053 | -35.5743
25.2655 | -35.5056 | -35.5797
25.2716 | -35.506 | -35.5851
25.2777 | -35.5063 | -35.5908
25.2838 | -35.5067 | -35.5964
25.29 | -35.5071 | -35.6019
25.2961 | -35.5074 | -35.6075
25.3022 | -35.5078 | -35.6131
25.3083 | -35.5081 | -35.6188
25.3144 | -35.5085 | -35.6245
25.3205 | -35.5089 | -35.63
25.3266 | -35.5092 | -35.6353
25.3327 | -35.5096 | -35.6413
25.3388 | -35.5099 | -35.6468
25.3449 | -35.5102 | -35.6524
25.351 | -35.5106 | -35.6579
25.3571 | -35.5109 | -35.6632
25.3632 | -35.5113 | -35.6687
25.3693 | -35.5117 | -35.6741
25.3754 | -35.512 | -35.6796
25.3815 | -35.5124 | -35.6852
25.3876 | -35.5128 | -35.6906
25.3937 | -35.5131 | -35.6959
25.3998 | -35.5135 | -35.7011
25.4059 | -35.5139 | -35.7066
25.412 | -35.5142 | -35.7119
25.4181 | -35.5146 | -35.7174
25.4242 | -35.515 | -35.7231
25.4304 | -35.5153 | -35.729
25.4365 | -35.5157 | -35.7345
25.4426 | -35.5162 | -35.7402
25.4487 | -35.5165 | -35.7457
25.4548 | -35.5168 | -35.7511
25.4609 | -35.5172 | -35.757
25.467 | -35.5176 | -35.7619
25.4731 | -35.518 | -35.7675
25.4792 | -35.5184 | -35.7727
25.4853 | -35.5188 | -35.7781
25.4914 | -35.5192 | -35.7836
25.4975 | -35.5195 | -35.7893
25.5036 | -35.5199 | -35.7946
25.5097 | -35.5203 | -35.8
25.5158 | -35.5207 | -35.8053
25.5219 | -35.5211 | -35.8108
25.528 | -35.5215 | -35.8159
25.5341 | -35.5219 | -35.8215
25.5402 | -35.5222 | -35.8266
25.5463 | -35.5225 | -35.8317
25.5524 | -35.523 | -35.8372
25.5585 | -35.5234 | -35.843
25.5646 | -35.5238 | -35.8488
25.5707 | -35.5241 | -35.8544
25.5769 | -35.5245 | -35.8599
25.583 | -35.5249 | -35.8655
25.5891 | -35.5253 | -35.8708
25.5952 | -35.5257 | -35.8762
25.6013 | -35.5261 | -35.8818
25.6074 | -35.5265 | -35.8873
25.6135 | -35.5269 | -35.8927
25.6196 | -35.5273 | -35.8985
25.6257 | -35.5277 | -35.9038
25.6318 | -35.528 | -35.9091
25.6379 | -35.5284 | -35.9148
25.644 | -35.5287 | -35.92
25.6501 | -35.529 | -35.9253
25.6562 | -35.5294 | -35.9307
25.6623 | -35.5298 | -35.9365
25.6684 | -35.5301 | -35.942
25.6745 | -35.5304 | -35.9479
25.6806 | -35.5307 | -35.9539
25.6867 | -35.531 | -35.96
25.6928 | -35.5313 | -35.9652
25.6989 | -35.5317 | -35.9706
25.705 | -35.5322 | -35.9765
25.7111 | -35.5326 | -35.9824
25.7173 | -35.533 | -35.9866
25.7234 | -35.5334 | -35.9922
25.7295 | -35.5337 | -35.9978
25.7356 | -35.5341 | -36.0031
25.7417 | -35.5345 | -36.0084
25.7478 | -35.5349 | -36.0138
25.7539 | -35.5353 | -36.019
25.76 | -35.5356 | -36.0247
25.7661 | -35.5359 | -36.0304
25.7722 | -35.5363 | -36.0358
25.7783 | -35.5366 | -36.0413
25.7844 | -35.537 | -36.0469
25.7905 | -35.5373 | -36.0522
25.7966 | -35.5378 | -36.0573
25.8027 | -35.5381 | -36.0632
25.8088 | -35.5385 | -36.069
25.8149 | -35.5388 | -36.0744
25.821 | -35.5391 | -36.0796
25.8271 | -35.5396 | -36.0853
25.8332 | -35.5399 | -36.0905
25.8393 | -35.5403 | -36.0961
25.8454 | -35.5406 | -36.1018
25.8515 | -35.541 | -36.1073
25.8576 | -35.5413 | -36.1128
25.8638 | -35.5416 | -36.1182
25.8699 | -35.5419 | -36.1237
25.876 | -35.5422 | -36.1292
25.8821 | -35.5426 | -36.1346
25.8882 | -35.5429 | -36.14
25.8943 | -35.5432 | -36.1458
25.9004 | -35.5436 | -36.1512
25.9065 | -35.544 | -36.1566
25.9126 | -35.5443 | -36.162
25.9187 | -35.5447 | -36.1673
25.9248 | -35.545 | -36.172
25.9309 | -35.5454 | -36.1773
25.937 | -35.5457 | -36.1829
25.9431 | -35.5461 | -36.1885
25.9492 | -35.5464 | -36.1936
25.9553 | -35.5467 | -36.1994
25.9614 | -35.547 | -36.2047
25.9675 | -35.5474 | -36.2104
25.9736 | -35.5478 | -36.2161
25.9797 | -35.5481 | -36.2222
25.9858 | -35.5486 | -36.2275
25.9919 | -35.5489 | -36.2331
25.998 | -35.5491 | -36.2387
26.0042 | -35.5495 | -36.2444
26.0103 | -35.5499 | -36.2501
26.0164 | -35.5503 | -36.2552
26.0225 | -35.5506 | -36.2609
26.0286 | -35.551 | -36.2658
26.0347 | -35.5514 | -36.2719
26.0408 | -35.5518 | -36.2776
26.0469 | -35.5521 | -36.2834
26.053 | -35.5525 | -36.2893
26.0591 | -35.5528 | -36.2945
26.0652 | -35.5531 | -36.2996
26.0713 | -35.5534 | -36.3049
26.0774 | -35.5538 | -36.3104
26.0835 | -35.5542 | -36.316
26.0896 | -35.5545 | -36.3213
26.0957 | -35.5548 | -36.3267
26.1018 | -35.5551 | -36.3322
26.1079 | -35.5555 | -36.3378
26.114 | -35.5558 | -36.343
26.1201 | -35.5561 | -36.3484
26.1262 | -35.5565 | -36.3537
26.1323 | -35.5568 | -36.3588
26.1384 | -35.5572 | -36.3644
26.1445 | -35.5576 | -36.3701
26.1507 | -35.5579 | -36.3753
26.1568 | -35.5583 | -36.3808
26.1629 | -35.5586 | -36.3864
26.169 | -35.559 | -36.392
26.1751 | -35.5593 | -36.398
26.1812 | -35.5596 | -36.4033
26.1873 | -35.56 | -36.4086
26.1934 | -35.5603 | -36.4135
26.1995 | -35.5607 | -36.4189
26.2056 | -35.561 | -36.4246
26.2117 | -35.5613 | -36.4301
26.2178 | -35.5617 | -36.4354
26.2239 | -35.562 | -36.4406
26.23 | -35.5624 | -36.446
26.2361 | -35.5627 | -36.4517
26.2422 | -35.5631 | -36.457
26.2483 | -35.5634 | -36.4622
26.2544 | -35.5638 | -36.4674
26.2605 | -35.5641 | -36.4731
26.2666 | -35.5645 | -36.4785
26.2727 | -35.5648 | -36.484
26.2788 | -35.5652 | -36.4891
26.2849 | -35.5656 | -36.4944
26.2911 | -35.5661 | -36.4998
26.2972 | -35.5664 | -36.5048
26.3033 | -35.5668 | -36.5103
26.3094 | -35.5672 | -36.5155
26.3155 | -35.5675 | -36.5206
26.3216 | -35.5678 | -36.526
26.3277 | -35.5682 | -36.5317
26.3338 | -35.5686 | -36.537
26.3399 | -35.569 | -36.5423
26.346 | -35.5692 | -36.5481
26.3521 | -35.5695 | -36.5537
26.3582 | -35.5698 | -36.5597
26.3643 | -35.5702 | -36.565
26.3704 | -35.5705 | -36.5701
26.3765 | -35.5709 | -36.5761
26.3826 | -35.5712 | -36.5811
26.3887 | -35.5715 | -36.5863
26.3948 | -35.5719 | -36.5915
26.4009 | -35.5722 | -36.5966
26.407 | -35.5725 | -36.6015
26.4131 | -35.5729 | -36.6074
26.4192 | -35.5733 | -36.6123
26.4253 | -35.5736 | -36.6179
26.4314 | -35.574 | -36.6233
26.4376 | -35.5744 | -36.6289
26.4437 | -35.5748 | -36.6343
26.4498 | -35.5751 | -36.6403
26.4559 | -35.5755 | -36.6454
26.462 | -35.5758 | -36.6508
26.4681 | -35.5762 | -36.6562
26.4742 | -35.5765 | -36.6617
26.4803 | -35.5769 | -36.6671
26.4864 | -35.5772 | -36.6723
26.4925 | -35.5776 | -36.6781
26.4986 | -35.5779 | -36.6837
26.5047 | -35.5782 | -36.6889
26.5108 | -35.5786 | -36.6942
26.5169 | -35.5789 | -36.6997
26.523 | -35.5791 | -36.7047
26.5291 | -35.5795 | -36.7099
26.5352 | -35.5798 | -36.7156
26.5413 | -35.5801 | -36.7214
26.5474 | -35.5804 | -36.7265
26.5535 | -35.5808 | -36.7316
26.5596 | -35.5812 | -36.7367
26.5657 | -35.5816 | -36.7423
26.5718 | -35.5819 | -36.7483
26.578 | -35.5822 | -36.7539
26.5841 | -35.5826 | -36.7592
26.5902 | -35.5829 | -36.7646
26.5963 | -35.5831 | -36.77
26.6024 | -35.5835 | -36.7756
26.6085 | -35.5838 | -36.7809
26.6146 | -35.5841 | -36.7862
26.6207 | -35.5845 | -36.7911
26.6268 | -35.5848 | -36.7965
26.6329 | -35.5851 | -36.8018
26.639 | -35.5855 | -36.8069
26.6451 | -35.5859 | -36.8125
26.6512 | -35.5862 | -36.8183
26.6573 | -35.5866 | -36.8231
26.6634 | -35.5869 | -36.8283
26.6695 | -35.5872 | -36.8343
26.6756 | -35.5876 | -36.8401
26.6817 | -35.588 | -36.8458
26.6878 | -35.5883 | -36.8512
26.6939 | -35.5886 | -36.8569
26.7 | -35.589 | -36.8619
26.7061 | -35.5893 | -36.8674
26.7122 | -35.5896 | -36.8725
26.7183 | -35.59 | -36.8779
26.7245 | -35.5903 | -36.8834
26.7306 | -35.5906 | -36.8886
26.7367 | -35.591 | -36.8939
26.7428 | -35.5914 | -36.8993
26.7489 | -35.5917 | -36.9044
26.755 | -35.5921 | -36.9099
26.7611 | -35.5924 | -36.9154
26.7672 | -35.5928 | -36.9207
26.7733 | -35.5931 | -36.9264
26.7794 | -35.5934 | -36.932
26.7855 | -35.5938 | -36.9376
26.7916 | -35.5941 | -36.9428
26.7977 | -35.5944 | -36.9481
26.8038 | -35.5948 | -36.9537
26.8099 | -35.5952 | -36.9591
26.816 | -35.5955 | -36.9645
26.8221 | -35.5959 | -36.9701
26.8282 | -35.5962 | -36.9755
26.8343 | -35.5965 | -36.9809
26.8404 | -35.5968 | -36.986
26.8465 | -35.5972 | -36.9912
26.8526 | -35.5974 | -36.9965
26.8587 | -35.5978 | -37.002
26.8649 | -35.5981 | -37.0073
26.871 | -35.5985 | -37.0131
26.8771 | -35.5988 | -37.0186
26.8832 | -35.5992 | -37.0242
26.8893 | -35.5995 | -37.0299
26.8954 | -35.5998 | -37.0353
26.9015 | -35.6001 | -37.0407
26.9076 | -35.6004 | -37.046
26.9137 | -35.6007 | -37.0512
26.9198 | -35.6011 | -37.0565
26.9259 | -35.6014 | -37.0617
26.932 | -35.6017 | -37.0668
26.9381 | -35.6021 | -37.0724
26.9442 | -35.6024 | -37.078
26.9503 | -35.6028 | -37.0831
26.9564 | -35.6031 | -37.0885
26.9625 | -35.6034 | -37.0939
26.9686 | -35.6037 | -37.0993
26.9747 | -35.6041 | -37.1048
26.9808 | -35.6044 | -37.1099
26.9869 | -35.6047 | -37.1151
26.993 | -35.605 | -37.1206
26.9991 | -35.6054 | -37.1258
27.0052 | -35.6057 | -37.1313
27.0114 | -35.606 | -37.137
27.0175 | -35.6062 | -37.1423
27.0236 | -35.6066 | -37.1481
27.0297 | -35.6069 | -37.1537
27.0358 | -35.6073 | -37.159
27.0419 | -35.6077 | -37.1644
27.048 | -35.608 | -37.1701
27.0541 | -35.6083 | -37.1754
27.0602 | -35.6086 | -37.1808
27.0663 | -35.6089 | -37.1863
27.0724 | -35.6093 | -37.1919
27.0785 | -35.6096 | -37.1974
27.0846 | -35.6099 | -37.203
27.0907 | -35.6102 | -37.2086
27.0968 | -35.6106 | -37.2136
27.1029 | -35.6109 | -37.219
27.109 | -35.6112 | -37.2244
27.1151 | -35.6116 | -37.2299
27.1212 | -35.6119 | -37.2351
27.1273 | -35.6122 | -37.2405
27.1334 | -35.6126 | -37.246
27.1395 | -35.6129 | -37.251
27.1456 | -35.6132 | -37.2564
27.1518 | -35.6135 | -37.2613
27.1579 | -35.6138 | -37.2662
27.164 | -35.6141 | -37.2714
27.1701 | -35.6145 | -37.2765
27.1762 | -35.6148 | -37.2819
27.1823 | -35.6151 | -37.2875
27.1884 | -35.6154 | -37.2928
27.1945 | -35.6158 | -37.2974
27.2006 | -35.6161 | -37.3024
27.2067 | -35.6164 | -37.3075
27.2128 | -35.6167 | -37.3128
27.2189 | -35.617 | -37.3181
27.225 | -35.6174 | -37.3236
27.2311 | -35.6178 | -37.3287
27.2372 | -35.6181 | -37.3337
27.2433 | -35.6184 | -37.3391
27.2494 | -35.6187 | -37.344
27.2555 | -35.619 | -37.3495
27.2616 | -35.6193 | -37.3549
27.2677 | -35.6195 | -37.36
27.2738 | -35.6198 | -37.3653
27.2799 | -35.6202 | -37.3707
27.286 | -35.6205 | -37.3761
27.2921 | -35.6209 | -37.3816
27.2983 | -35.6212 | -37.3869
27.3044 | -35.6215 | -37.3922
27.3105 | -35.6218 | -37.3974
27.3166 | -35.6222 | -37.4028
27.3227 | -35.6225 | -37.4082
27.3288 | -35.6228 | -37.4134
27.3349 | -35.6232 | -37.4183
27.341 | -35.6234 | -37.4237
27.3471 | -35.6238 | -37.4285
27.3532 | -35.6241 | -37.434
27.3593 | -35.6245 | -37.4392
27.3654 | -35.6248 | -37.4447
27.3715 | -35.6252 | -37.4498
27.3776 | -35.6255 | -37.4552
27.3837 | -35.6258 | -37.4607
27.3898 | -35.6261 | -37.4659
27.3959 | -35.6265 | -37.4713
27.402 | -35.6268 | -37.4768
27.4081 | -35.6271 | -37.4822
27.4142 | -35.6274 | -37.4874
27.4203 | -35.6277 | -37.4928
27.4264 | -35.628 | -37.498
27.4325 | -35.6284 | -37.5034
27.4387 | -35.6287 | -37.5087
27.4448 | -35.629 | -37.514
27.4509 | -35.6292 | -37.5194
27.457 | -35.6296 | -37.5245
27.4631 | -35.6299 | -37.5299
27.4692 | -35.6302 | -37.5351
27.4753 | -35.6305 | -37.5403
27.4814 | -35.6308 | -37.5458
27.4875 | -35.6311 | -37.5509
27.4936 | -35.6314 | -37.5561
27.4997 | -35.6317 | -37.5615
27.5058 | -35.632 | -37.5671
27.5119 | -35.6323 | -37.5723
27.518 | -35.6327 | -37.5776
27.5241 | -35.633 | -37.5831
27.5302 | -35.6333 | -37.5883
27.5363 | -35.6336 | -37.5934
27.5424 | -35.6339 | -37.5983
27.5485 | -35.6343 | -37.6038
27.5546 | -35.6346 | -37.609
27.5607 | -35.6349 | -37.6142
27.5668 | -35.6351 | -37.6194
27.5729 | -35.6354 | -37.6245
27.5791 | -35.6358 | -37.6297
27.5852 | -35.6361 | -37.6349
27.5913 | -35.6365 | -37.6401
27.5974 | -35.6368 | -37.6454
27.6035 | -35.6371 | -37.6509
27.6096 | -35.6374 | -37.6562
27.6157 | -35.6377 | -37.6614
27.6218 | -35.6381 | -37.6667
27.6279 | -35.6384 | -37.6718
27.634 | -35.6387 | -37.677
27.6401 | -35.639 | -37.6821
27.6462 | -35.6394 | -37.6874
27.6523 | -35.6397 | -37.6924
27.6584 | -35.64 | -37.6981
27.6645 | -35.6403 | -37.7032
27.6706 | -35.6406 | -37.7087
27.6767 | -35.6409 | -37.7138
27.6828 | -35.6412 | -37.7193
27.6889 | -35.6415 | -37.7246
27.695 | -35.6418 | -37.7299
27.7011 | -35.6421 | -37.7354
27.7072 | -35.6425 | -37.7408
27.7133 | -35.6428 | -37.7459
27.7194 | -35.643 | -37.7513
27.7256 | -35.6434 | -37.7565
27.7317 | -35.6437 | -37.7617
27.7378 | -35.644 | -37.767
27.7439 | -35.6443 | -37.7724
27.75 | -35.6446 | -37.7779
27.7561 | -35.645 | -37.783
27.7622 | -35.6453 | -37.7883
27.7683 | -35.6456 | -37.7935
27.7744 | -35.646 | -37.7987
27.7805 | -35.6463 | -37.804
27.7866 | -35.6466 | -37.8093
27.7927 | -35.6469 | -37.8141
27.7988 | -35.6472 | -37.8193
27.8049 | -35.6475 | -37.8248
27.811 | -35.6478 | -37.8302
27.8171 | -35.6481 | -37.8354
27.8232 | -35.6485 | -37.8403
27.8293 | -35.6488 | -37.8454
27.8354 | -35.6491 | -37.8507
27.8415 | -35.6494 | -37.8558
27.8476 | -35.6498 | -37.861
27.8537 | -35.6501 | -37.8661
27.8598 | -35.6503 | -37.8716
27.866 | -35.6507 | -37.8767
27.8721 | -35.651 | -37.8817
27.8782 | -35.6513 | -37.8871
27.8843 | -35.6516 | -37.8922
27.8904 | -35.6519 | -37.8973
27.8965 | -35.6522 | -37.9024
27.9026 | -35.6526 | -37.9079
27.9087 | -35.6529 | -37.9132
27.9148 | -35.6532 | -37.9186
27.9209 | -35.6535 | -37.924
27.927 | -35.6538 | -37.929
27.9331 | -35.6541 | -37.9343
27.9392 | -35.6544 | -37.9394
27.9453 | -35.6547 | -37.9446
27.9514 | -35.655 | -37.9499
27.9575 | -35.6553 | -37.9554
27.9636 | -35.6555 | -37.9607
27.9697 | -35.6558 | -37.966
27.9758 | -35.6561 | -37.9713
27.9819 | -35.6564 | -37.9766
27.988 | -35.6567 | -37.9818
27.9941 | -35.657 | -37.9871
28.0002 | -35.6573 | -37.9922
28.0063 | -35.6576 | -37.9973
28.0125 | -35.6579 | -38.0026
28.0186 | -35.6583 | -38.0077
28.0247 | -35.6586 | -38.013
28.0308 | -35.6589 | -38.0184
28.0369 | -35.6592 | -38.0236
28.043 | -35.6595 | -38.0289
28.0491 | -35.6598 | -38.0341
28.0552 | -35.6601 | -38.0394
28.0613 | -35.6604 | -38.0443
28.0674 | -35.6607 | -38.0494
28.0735 | -35.661 | -38.0544

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5475 | -42.2319
21.4908 | -35.5561 | -35.1591
16.4948 | -34.9321 | -27.7657
12.6602 | -34.5723 | -21.4857
9.7171 | -34.3524 | -16.5614
7.4582 | -34.2039 | -12.5838
5.7244 | -34.097 | -9.3494
4.3936 | -34.0228 | -6.7488
3.3722 | -34.0087 | -4.6923
2.5883 | -33.9891 | -3.1336
1.9866 | -33.9899 | -1.8916
1.5248 | -33.9868 | -0.89309
1.1703 | -33.9846 | -0.24034
1 | -33.9869 | 0.079403
0.89822 | -33.9789 | 0.3535
0.68941 | -33.9673 | 0.93506
0.52919 | -33.9757 | 1.3864
0.40612 | -33.9648 | 1.9934
0.31171 | -33.9504 | 2.4535
0.23928 | -33.969 | 3.1747
0.18358 | -33.9568 | 3.5939
0.14091 | -33.9815 | 4.4728
0.10819 | -33.985 | 5.6385
0.083017 | -34.0234 | 7.1124
0.063705 | -34.0762 | 8.9801
0.048923 | -34.1389 | 11.4736
0.037479 | -34.2414 | 14.6614
0.028801 | -34.4324 | 18.6674
0.022078 | -34.7508 | 23.922
0.017023 | -35.226 | 30.2932
0.013018 | -36.0181 | 37.9678
0.010014 | -37.1313 | 45.8916

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003643 | -4.5703
0.007286 | -29.3005
0.010929 | -37.3212
0.014572 | -42.1461
0.018215 | -44.1684
0.021858 | -47.0027
0.025501 | -47.8562
0.029144 | -48.4486
0.032787 | -49.5264
0.03643 | -52.1769
0.040073 | -53.3163
0.043716 | -55.2101
0.047359 | -55.7041
0.051002 | -56.7728
0.054645 | -58.1053
0.058288 | -58.0134
0.061931 | -58.6727
0.065574 | -59.7834
0.069217 | -60.1477
0.07286 | -61.6906
0.076503 | -61.3007
0.080146 | -60.3829
0.083789 | -61.3649
0.087432 | -62.4813
0.091075 | -63.6644
0.094718 | -63.984
0.098361 | -63.6436
0.102 | -63.9489
0.10565 | -64.2032
0.10929 | -64.6234
0.11293 | -65.4536
0.11658 | -65.8371
0.12022 | -65.5507
0.12386 | -66.7758
0.12751 | -67.5657
0.13115 | -68.2642
0.13479 | -67.7034
0.13843 | -66.7595
0.14208 | -66.9748
0.14572 | -68.1298
0.14936 | -69.2591
0.15301 | -69.3196
0.15665 | -69.4445
0.16029 | -69.6545
0.16393 | -69.4773
0.16758 | -69.7827
0.17122 | -69.6128
0.17486 | -70.0055
0.17851 | -72.2373
0.18215 | -71.9373
0.18579 | -71.7523
0.18943 | -71.5393
0.19308 | -71.0966
0.19672 | -71.5642
0.20036 | -72.14
0.20401 | -72.4949
0.20765 | -73.2895
0.21129 | -73.5711
0.21494 | -73.3662
0.21858 | -73.7089
0.22222 | -73.7276
0.22587 | -73.3151
0.22951 | -73.2042
0.23315 | -73.4684
0.23679 | -73.1843
0.24044 | -74.824
0.24408 | -74.8303
0.24772 | -74.141
0.25137 | -74.3195
0.25501 | -74.1983
0.25865 | -74.495
0.2623 | -75.3896
0.26594 | -74.8369
0.26958 | -74.962
0.27322 | -76.0111
0.27687 | -75.6256
0.28051 | -75.395
0.28415 | -76.14
0.2878 | -75.6135
0.29144 | -75.6138
0.29508 | -75.198
0.29873 | -74.9651
0.30237 | -75.9219
0.30601 | -75.3919
0.30965 | -74.3993
0.3133 | -75.6694
0.31694 | -76.8049
0.32058 | -76.3389
0.32423 | -75.5321
0.32787 | -76.173
0.33151 | -76.5008
0.33515 | -76.8505
0.3388 | -77.0945
0.34244 | -75.2817
0.34608 | -75.4022
0.34973 | -77.0864
0.35337 | -76.762
0.35701 | -76.8919
0.36066 | -77.4884
0.3643 | -77.6851
0.36794 | -77.9586
0.37158 | -78.202
0.37523 | -78.5736
0.37887 | -77.7037
0.38251 | -78.805
0.38616 | -79.3336
0.3898 | -79.2342
0.39344 | -79.4883
0.39709 | -79.4155
0.40073 | -79.6511
0.40437 | -79.135
0.40802 | -79.1817
0.41166 | -79.6532
0.4153 | -79.9591
0.41894 | -79.2768
0.42259 | -79.6214
0.42623 | -79.5535
0.42987 | -79.1734
0.43351 | -78.6926
0.43716 | -78.9312
0.4408 | -78.7731
0.44444 | -78.4389
0.44809 | -78.8429
0.45173 | -79.3878
0.45537 | -80.3019
0.45902 | -79.5563
0.46266 | -79.7459
0.4663 | -79.6353
0.46995 | -79.4249
0.47359 | -80.2139
0.47723 | -81.1987
0.48087 | -81.1797
0.48452 | -81.2153
0.48816 | -80.4562
0.4918 | -80.0612
0.49545 | -80.4299
0.49909 | -80.908
0.50273 | -80.9857
0.50638 | -80.7181
0.51002 | -80.7839
0.51366 | -81.9138
0.5173 | -81.5012
0.52095 | -80.969
0.52459 | -80.509
0.52823 | -80.8693
0.53188 | -81.4626
0.53552 | -81.654
0.53916 | -81.8491
0.5428 | -82.37
0.54645 | -82.4843
0.55009 | -82.1909
0.55373 | -82.0167
0.55738 | -81.7121
0.56102 | -81.407
0.56466 | -81.3508
0.56831 | -81.4798
0.57195 | -81.7076
0.57559 | -82.1617
0.57924 | -82.4392
0.58288 | -82.4262
0.58652 | -82.2042
0.59016 | -81.912
0.59381 | -81.6938
0.59745 | -81.8037
0.60109 | -82.2675
0.60474 | -82.704
0.60838 | -82.8616
0.61202 | -83.2114
0.61567 | -83.5423
0.61931 | -83.5576
0.62295 | -83.0837
0.62659 | -82.6391
0.63024 | -83.0174
0.63388 | -83.6804
0.63752 | -83.8974
0.64117 | -83.6173
0.64481 | -83.2192
0.64845 | -83.0605
0.65209 | -83.0728
0.65574 | -83.2258
0.65938 | -83.3921
0.66302 | -83.3975
0.66667 | -83.2061
0.67031 | -83.4833
0.67395 | -83.8927
0.6776 | -83.9154
0.68124 | -83.996
0.68488 | -84.1664
0.68853 | -84.0745
0.69217 | -84.1081
0.69581 | -84.4095
0.69945 | -84.8508
0.7031 | -85.0341
0.70674 | -85.0029
0.71038 | -84.6006
0.71403 | -84.0925
0.71767 | -83.9427
0.72131 | -84.3594
0.72495 | -85.0552
0.7286 | -85.1234
0.73224 | -84.4342
0.73588 | -84.142
0.73953 | -84.5521
0.74317 | -84.6728
0.74681 | -84.2843
0.75045 | -84.2329
0.7541 | -84.829
0.75774 | -85.3312
0.76138 | -84.7662
0.76503 | -84.1861
0.76867 | -84.6443
0.77231 | -85.4572
0.77596 | -85.4887
0.7796 | -85.1377
0.78324 | -85.1367
0.78688 | -85.2991
0.79053 | -85.1994
0.79417 | -85.0153
0.79781 | -85.1296
0.80146 | -85.6882
0.8051 | -86.0648
0.80874 | -85.6167
0.81239 | -85.1478
0.81603 | -85.3815
0.81967 | -85.7997
0.82332 | -85.8467
0.82696 | -85.5669
0.8306 | -85.5496
0.83424 | -85.7543
0.83789 | -85.7662
0.84153 | -85.4332
0.84517 | -85.6834
0.84882 | -86.091
0.85246 | -85.8847
0.8561 | -85.9941
0.85974 | -86.6923
0.86339 | -87.2077
0.86703 | -87.022
0.87067 | -87.0278
0.87432 | -87.2588
0.87796 | -87.3324
0.8816 | -87.5128
0.88525 | -87.7096
0.88889 | -87.2344
0.89253 | -86.4771
0.89618 | -85.6438
0.89982 | -85.3796
0.90346 | -85.6972
0.9071 | -85.9905
0.91075 | -86.1584
0.91439 | -86.3056
0.91803 | -86.2956
0.92168 | -86.3285
0.92532 | -86.7862
0.92896 | -87.0663
0.93261 | -87.0129
0.93625 | -87.178
0.93989 | -87.7906
0.94353 | -88.1078
0.94718 | -87.8273
0.95082 | -87.1509
0.95446 | -86.7997
0.95811 | -86.6238
0.96175 | -86.7246
0.96539 | -87.1389
0.96903 | -87.4125
0.97268 | -87.2971
0.97632 | -86.9888
0.97996 | -86.9132
0.98361 | -86.9529
0.98725 | -86.803
0.99089 | -86.843
0.99454 | -86.9059
0.99818 | -86.6206
1.0018 | -86.7339
1.0055 | -87.4054
1.0091 | -87.836
1.0128 | -88.0613
1.0164 | -88.3411
1.02 | -88.7114
1.0237 | -88.8552
1.0273 | -88.2661
1.031 | -87.6123
1.0346 | -87.5525
1.0383 | -87.9053
1.0419 | -87.9038
1.0455 | -87.7684
1.0492 | -87.966
1.0528 | -88.0183
1.0565 | -87.9533
1.0601 | -87.8819
1.0638 | -87.8691
1.0674 | -87.8995
1.071 | -87.9537
1.0747 | -87.9993
1.0783 | -88.0619
1.082 | -88.183
1.0856 | -88.2569
1.0893 | -88.227
1.0929 | -88.17
1.0965 | -88.1074
1.1002 | -88.1264
1.1038 | -88.3234
1.1075 | -88.5335
1.1111 | -88.6051
1.1148 | -88.6179
1.1184 | -88.5657
1.122 | -88.3953
1.1257 | -88.2784
1.1293 | -88.4208
1.133 | -88.8039
1.1366 | -89.1499
1.1403 | -89.2788
1.1439 | -89.3937
1.1475 | -89.6593
1.1512 | -89.8434
1.1548 | -89.6641
1.1585 | -89.2706
1.1621 | -88.9182
1.1658 | -88.6987
1.1694 | -88.6882
1.173 | -88.8063
1.1767 | -88.822
1.1803 | -88.6963
1.184 | -88.5252
1.1876 | -88.363
1.1913 | -88.3623
1.1949 | -88.6275
1.1985 | -89.0372
1.2022 | -89.3689
1.2058 | -89.4895
1.2095 | -89.4543
1.2131 | -89.418
1.2168 | -89.4674
1.2204 | -89.6125
1.224 | -89.7163
1.2277 | -89.6408
1.2313 | -89.4529
1.235 | -89.2918
1.2386 | -89.294
1.2423 | -89.522
1.2459 | -89.7885
1.2495 | -89.784
1.2532 | -89.4284
1.2568 | -88.9763
1.2605 | -88.7029
1.2641 | -88.7831
1.2678 | -89.1913
1.2714 | -89.6315
1.275 | -89.8227
1.2787 | -89.8096
1.2823 | -89.7368
1.286 | -89.7412
1.2896 | -89.9414
1.2933 | -90.2285
1.2969 | -90.3824
1.3005 | -90.4742
1.3042 | -90.5841
1.3078 | -90.5992
1.3115 | -90.436
1.3151 | -90.1551
1.3188 | -89.906
1.3224 | -89.7662
1.326 | -89.7979
1.3297 | -90.0664
1.3333 | -90.5058
1.337 | -90.8862
1.3406 | -90.9694
1.3443 | -90.7327
1.3479 | -90.4555
1.3515 | -90.3723
1.3552 | -90.4604
1.3588 | -90.5512
1.3625 | -90.5954
1.3661 | -90.6055
1.3698 | -90.5261
1.3734 | -90.4322
1.377 | -90.4033
1.3807 | -90.337
1.3843 | -90.2618
1.388 | -90.3388
1.3916 | -90.5369
1.3953 | -90.7064
1.3989 | -90.8441
1.4025 | -90.9535
1.4062 | -90.9677
1.4098 | -90.9148
1.4135 | -90.8705
1.4171 | -90.8552
1.4208 | -90.7432
1.4244 | -90.5946
1.4281 | -90.6898
1.4317 | -91.0058
1.4353 | -91.2312
1.439 | -91.2369
1.4426 | -91.1702
1.4463 | -91.1072
1.4499 | -90.976
1.4536 | -90.797
1.4572 | -90.7562
1.4608 | -90.9486
1.4645 | -91.176
1.4681 | -91.2604
1.4718 | -91.2862
1.4754 | -91.3607
1.4791 | -91.4352
1.4827 | -91.434
1.4863 | -91.5263
1.49 | -91.8498
1.4936 | -92.1801
1.4973 | -92.1841
1.5009 | -91.9336
1.5046 | -91.7936
1.5082 | -91.9075
1.5118 | -92.1387
1.5155 | -92.3224
1.5191 | -92.3851
1.5228 | -92.28
1.5264 | -91.9786
1.5301 | -91.6689
1.5337 | -91.5811
1.5373 | -91.7238
1.541 | -91.9368
1.5446 | -92.0839
1.5483 | -92.0051
1.5519 | -91.6501
1.5556 | -91.269
1.5592 | -91.1028
1.5628 | -91.1669
1.5665 | -91.3107
1.5701 | -91.3865
1.5738 | -91.3785
1.5774 | -91.375
1.5811 | -91.3729
1.5847 | -91.4195
1.5883 | -91.5717
1.592 | -91.8216
1.5956 | -92.0704
1.5993 | -92.158
1.6029 | -91.9748
1.6066 | -91.5866
1.6102 | -91.197
1.6138 | -90.9835
1.6175 | -90.9871
1.6211 | -91.1202
1.6248 | -91.288
1.6284 | -91.4572
1.6321 | -91.6454
1.6357 | -91.8666
1.6393 | -92.0896
1.643 | -92.2205
1.6466 | -92.1981
1.6503 | -92.0694
1.6539 | -91.887
1.6576 | -91.7018
1.6612 | -91.5792
1.6648 | -91.5556
1.6685 | -91.6169
1.6721 | -91.7474
1.6758 | -91.9567
1.6794 | -92.2044
1.6831 | -92.4611
1.6867 | -92.656
1.6903 | -92.6423
1.694 | -92.4163
1.6976 | -92.1725
1.7013 | -92.074
1.7049 | -92.1079
1.7086 | -92.1665
1.7122 | -92.1817
1.7158 | -92.1749
1.7195 | -92.1788
1.7231 | -92.1886
1.7268 | -92.2224
1.7304 | -92.3164
1.7341 | -92.4352
1.7377 | -92.5049
1.7413 | -92.4884
1.745 | -92.4423
1.7486 | -92.4218
1.7523 | -92.44
1.7559 | -92.4627
1.7596 | -92.4374
1.7632 | -92.3599
1.7668 | -92.2682
1.7705 | -92.1839
1.7741 | -92.148
1.7778 | -92.182
1.7814 | -92.2843
1.7851 | -92.4383
1.7887 | -92.6066
1.7924 | -92.7857
1.796 | -92.9451
1.7996 | -92.9918
1.8033 | -92.8609
1.8069 | -92.5742
1.8106 | -92.2719
1.8142 | -92.111
1.8179 | -92.1634
1.8215 | -92.359
1.8251 | -92.5013
1.8288 | -92.4301
1.8324 | -92.2259
1.8361 | -92.0773
1.8397 | -92.0763
1.8434 | -92.1751
1.847 | -92.2972
1.8506 | -92.4292
1.8543 | -92.5902
1.8579 | -92.754
1.8616 | -92.8803
1.8652 | -92.9729
1.8689 | -93.056
1.8725 | -93.1672
1.8761 | -93.2981
1.8798 | -93.4413
1.8834 | -93.5825
1.8871 | -93.6689
1.8907 | -93.5746
1.8944 | -93.2746
1.898 | -92.9618
1.9016 | -92.8456
1.9053 | -92.9357
1.9089 | -93.1236
1.9126 | -93.2735
1.9162 | -93.3085
1.9199 | -93.2238
1.9235 | -93.0751
1.9271 | -92.9447
1.9308 | -92.8867
1.9344 | -92.9296
1.9381 | -93.0809
1.9417 | -93.2728
1.9454 | -93.3527
1.949 | -93.2303
1.9526 | -93.0129
1.9563 | -92.8848
1.9599 | -92.9166
1.9636 | -93.041
1.9672 | -93.1511
1.9709 | -93.2175
1.9745 | -93.2997
1.9781 | -93.4285
1.9818 | -93.5737
1.9854 | -93.7178
1.9891 | -93.8927
1.9927 | -94.0621
1.9964 | -94.0871
2 | -93.8729
2.0036 | -93.5347
2.0073 | -93.2925
2.0109 | -93.2472
2.0146 | -93.3235
2.0182 | -93.3933
2.0219 | -93.4059
2.0255 | -93.4355
2.0291 | -93.5486
2.0328 | -93.7175
2.0364 | -93.8594
2.0401 | -93.9575
2.0437 | -94.0778
2.0474 | -94.249
2.051 | -94.4256
2.0546 | -94.5224
2.0583 | -94.5233
2.0619 | -94.4719
2.0656 | -94.371
2.0692 | -94.1745
2.0729 | -93.9142
2.0765 | -93.6829
2.0801 | -93.5782
2.0838 | -93.5977
2.0874 | -93.6642
2.0911 | -93.633
2.0947 | -93.5125
2.0984 | -93.4511
2.102 | -93.583
2.1056 | -93.6774
2.1093 | -93.7675
2.1129 | -93.8363
2.1166 | -93.8837
2.1202 | -93.9239
2.1239 | -93.9612
2.1275 | -94.0332
2.1311 | -94.1594
2.1348 | -94.3336
2.1384 | -94.5092
2.1421 | -94.6417
2.1457 | -94.7126
2.1494 | -94.7125
2.153 | -94.6716
2.1566 | -94.6078
2.1603 | -94.5627
2.1639 | -94.556
2.1676 | -94.5898
2.1712 | -94.6322
2.1749 | -94.6541
2.1785 | -94.623
2.1821 | -94.5255
2.1858 | -94.4061
2.1894 | -94.3419
2.1931 | -94.3827
2.1967 | -94.5488
2.2004 | -94.7844
2.204 | -94.9994
2.2077 | -95.0979
2.2113 | -95.0589
2.2149 | -94.8965
2.2186 | -94.6803
2.2222 | -94.459
2.2259 | -94.2542
2.2295 | -94.089
2.2332 | -93.9899
2.2368 | -93.9531
2.2404 | -93.9639
2.2441 | -94.0149
2.2477 | -94.0793
2.2514 | -94.1216
2.255 | -94.1421
2.2587 | -94.1483
2.2623 | -94.1727
2.2659 | -94.2605
2.2696 | -94.4343
2.2732 | -94.678
2.2769 | -94.9587
2.2805 | -95.2041
2.2842 | -95.3568
2.2878 | -95.4232
2.2914 | -95.4223
2.2951 | -95.3537
2.2987 | -95.2287
2.3024 | -95.0654
2.306 | -94.9093
2.3097 | -94.7985
2.3133 | -94.752
2.3169 | -94.722
2.3206 | -94.6808
2.3242 | -94.6519
2.3279 | -94.6802
2.3315 | -94.8012
2.3352 | -94.9864
2.3388 | -95.1716
2.3424 | -95.2986
2.3461 | -95.332
2.3497 | -95.3112
2.3534 | -95.2699
2.357 | -95.2368
2.3607 | -95.1883
2.3643 | -95.0913
2.3679 | -94.9717
2.3716 | -94.8522
2.3752 | -94.7359
2.3789 | -94.6193
2.3825 | -94.4851
2.3862 | -94.3438
2.3898 | -94.2285
2.3934 | -94.1733
2.3971 | -94.1916
2.4007 | -94.2576
2.4044 | -94.3187
2.408 | -94.3481
2.4117 | -94.3623
2.4153 | -94.4041
2.4189 | -94.5007
2.4226 | -94.6683
2.4262 | -94.8724
2.4299 | -95.0544
2.4335 | -95.1759
2.4372 | -95.2127
2.4408 | -95.1795
2.4444 | -95.0938
2.4481 | -94.9804
2.4517 | -94.8708
2.4554 | -94.7767
2.459 | -94.6928
2.4627 | -94.6313
2.4663 | -94.6355
2.4699 | -94.7465
2.4736 | -94.9594
2.4772 | -95.2115
2.4809 | -95.4091
2.4845 | -95.5034
2.4882 | -95.482
2.4918 | -95.3878
2.4954 | -95.2586
2.4991 | -95.095
2.5027 | -94.8999
2.5064 | -94.7013
2.51 | -94.5391
2.5137 | -94.4561
2.5173 | -94.4796
2.5209 | -94.6119
2.5246 | -94.8246
2.5282 | -95.0617
2.5319 | -95.2644
2.5355 | -95.4136
2.5392 | -95.5314
2.5428 | -95.6039
2.5464 | -95.6245
2.5501 | -95.5786
2.5537 | -95.4722
2.5574 | -95.356
2.561 | -95.2933
2.5647 | -95.3305
2.5683 | -95.4435
2.5719 | -95.559
2.5756 | -95.6183
2.5792 | -95.6324
2.5829 | -95.6542
2.5865 | -95.7011
2.5902 | -95.7766
2.5938 | -95.8538
2.5974 | -95.9109
2.6011 | -95.9468
2.6047 | -95.9512
2.6084 | -95.9426
2.612 | -95.9129
2.6157 | -95.867
2.6193 | -95.8479
2.623 | -95.8918
2.6266 | -96.0431
2.6302 | -96.1659
2.6339 | -96.2873
2.6375 | -96.3717
2.6412 | -96.3839
2.6448 | -96.3226
2.6485 | -96.1823
2.6521 | -95.9902
2.6557 | -95.7954
2.6594 | -95.6494
2.663 | -95.5658
2.6667 | -95.5623
2.6703 | -95.613
2.674 | -95.6804
2.6776 | -95.7446
2.6812 | -95.7872
2.6849 | -95.805
2.6885 | -95.7988
2.6922 | -95.7669
2.6958 | -95.73
2.6995 | -95.6939
2.7031 | -95.6728
2.7067 | -95.6718
2.7104 | -95.6823
2.714 | -95.6928
2.7177 | -95.6979
2.7213 | -95.6849
2.725 | -95.6511
2.7286 | -95.6169
2.7322 | -95.5947
2.7359 | -95.6063
2.7395 | -95.6425
2.7432 | -95.6876
2.7468 | -95.735
2.7505 | -95.7972
2.7541 | -95.8975
2.7577 | -96.0513
2.7614 | -96.2122
2.765 | -96.3158
2.7687 | -96.3128
2.7723 | -96.1867
2.776 | -96.0055
2.7796 | -95.8316
2.7832 | -95.7031
2.7869 | -95.6611
2.7905 | -95.7073
2.7942 | -95.8356
2.7978 | -96.0201
2.8015 | -96.2085
2.8051 | -96.3513
2.8087 | -96.4302
2.8124 | -96.4355
2.816 | -96.3782
2.8197 | -96.3102
2.8233 | -96.2509
2.827 | -96.2268
2.8306 | -96.2395
2.8342 | -96.2882
2.8379 | -96.3709
2.8415 | -96.464
2.8452 | -96.5537
2.8488 | -96.6379
2.8525 | -96.7024
2.8561 | -96.7206
2.8597 | -96.6557
2.8634 | -96.4984
2.867 | -96.2895
2.8707 | -96.101
2.8743 | -95.9983
2.878 | -95.9935
2.8816 | -96.0517
2.8852 | -96.141
2.8889 | -96.2407
2.8925 | -96.3373
2.8962 | -96.4055
2.8998 | -96.4462
2.9035 | -96.4639
2.9071 | -96.4841
2.9107 | -96.5295
2.9144 | -96.5996
2.918 | -96.6908
2.9217 | -96.7978
2.9253 | -96.9054
2.929 | -97.0023
2.9326 | -97.0753
2.9362 | -97.1178
2.9399 | -97.1264
2.9435 | -97.1124
2.9472 | -97.0688
2.9508 | -97.0113
2.9545 | -96.9569
2.9581 | -96.926
2.9617 | -96.9241
2.9654 | -96.9311
2.969 | -96.9216
2.9727 | -96.8886
2.9763 | -96.8186
2.98 | -96.7179
2.9836 | -96.5952
2.9872 | -96.4526
2.9909 | -96.3008
2.9945 | -96.1562
2.9982 | -96.0384
3.0018 | -95.959
3.0055 | -95.9133
3.0091 | -95.9009
3.0128 | -95.9127
3.0164 | -95.9402
3.02 | -95.9955
3.0237 | -96.0915
3.0273 | -96.2316
3.031 | -96.4184
3.0346 | -96.6366
3.0383 | -96.8514
3.0419 | -97.0019
3.0455 | -97.0409
3.0492 | -96.9895
3.0528 | -96.882
3.0565 | -96.7778
3.0601 | -96.7113
3.0638 | -96.6656
3.0674 | -96.6353
3.071 | -96.614
3.0747 | -96.5904
3.0783 | -96.5757
3.082 | -96.5668
3.0856 | -96.5717
3.0893 | -96.5955
3.0929 | -96.6328
3.0965 | -96.6785
3.1002 | -96.7381
3.1038 | -96.808
3.1075 | -96.898
3.1111 | -97.0072
3.1148 | -97.1086
3.1184 | -97.1949
3.122 | -97.2484
3.1257 | -97.2565
3.1293 | -97.2253
3.133 | -97.1617
3.1366 | -97.0994
3.1403 | -97.0622
3.1439 | -97.0597
3.1475 | -97.0922
3.1512 | -97.109
3.1548 | -97.1101
3.1585 | -97.0937
3.1621 | -97.0673
3.1658 | -97.0392
3.1694 | -97.0273
3.173 | -97.0532
3.1767 | -97.1073
3.1803 | -97.1891
3.184 | -97.2895
3.1876 | -97.4085
3.1913 | -97.536
3.1949 | -97.6572
3.1985 | -97.7516
3.2022 | -97.8071
3.2058 | -97.8203
3.2095 | -97.7826
3.2131 | -97.7092
3.2168 | -97.627
3.2204 | -97.5518
3.224 | -97.5005
3.2277 | -97.4831
3.2313 | -97.5161
3.235 | -97.5781
3.2386 | -97.6417
3.2423 | -97.6834
3.2459 | -97.6889
3.2495 | -97.6407
3.2532 | -97.5681
3.2568 | -97.4871
3.2605 | -97.4083
3.2641 | -97.3479
3.2678 | -97.2951
3.2714 | -97.2452
3.275 | -97.2003
3.2787 | -97.1641
3.2823 | -97.1372
3.286 | -97.1232
3.2896 | -97.1216
3.2933 | -97.1242
3.2969 | -97.1229
3.3005 | -97.1146
3.3042 | -97.1054
3.3078 | -97.0862
3.3115 | -97.0721
3.3151 | -97.079
3.3188 | -97.1108
3.3224 | -97.1778
3.326 | -97.2729
3.3297 | -97.3824
3.3333 | -97.5026
3.337 | -97.6009
3.3406 | -97.6489
3.3443 | -97.6529
3.3479 | -97.6047
3.3515 | -97.5113
3.3552 | -97.4059
3.3588 | -97.314
3.3625 | -97.2526
3.3661 | -97.2287
3.3698 | -97.2384
3.3734 | -97.275
3.377 | -97.3366
3.3807 | -97.4113
3.3843 | -97.4922
3.388 | -97.5587
3.3916 | -97.6087
3.3953 | -97.624
3.3989 | -97.5992
3.4026 | -97.5492
3.4062 | -97.4772
3.4098 | -97.3914
3.4135 | -97.3161
3.4171 | -97.2721
3.4208 | -97.2809
3.4244 | -97.3415
3.4281 | -97.4403
3.4317 | -97.5535
3.4353 | -97.6748
3.439 | -97.7853
3.4426 | -97.872
3.4463 | -97.9299
3.4499 | -97.9619
3.4536 | -97.9729
3.4572 | -97.9708
3.4608 | -97.9698
3.4645 | -97.9685
3.4681 | -97.9768
3.4718 | -97.9934
3.4754 | -98.0075
3.4791 | -98.011
3.4827 | -98.012
3.4863 | -98.0216
3.49 | -98.0366
3.4936 | -98.0602
3.4973 | -98.1052
3.5009 | -98.1657
3.5046 | -98.2295
3.5082 | -98.2985
3.5118 | -98.3744
3.5155 | -98.4463
3.5191 | -98.5079
3.5228 | -98.5365
3.5264 | -98.5222
3.5301 | -98.4824
3.5337 | -98.4383
3.5373 | -98.3927
3.541 | -98.3637
3.5446 | -98.365
3.5483 | -98.4028
3.5519 | -98.4562
3.5556 | -98.5247
3.5592 | -98.5986
3.5628 | -98.6661
3.5665 | -98.7089
3.5701 | -98.7268
3.5738 | -98.7169
3.5774 | -98.6797
3.5811 | -98.6237
3.5847 | -98.5495
3.5883 | -98.4649
3.592 | -98.3886
3.5956 | -98.3275
3.5993 | -98.2899
3.6029 | -98.2715
3.6066 | -98.2607
3.6102 | -98.248
3.6138 | -98.2379
3.6175 | -98.2327
3.6211 | -98.2346
3.6248 | -98.2376
3.6284 | -98.2337
3.6321 | -98.2386
3.6357 | -98.2495
3.6393 | -98.2771
3.643 | -98.3309
3.6466 | -98.3968
3.6503 | -98.4676
3.6539 | -98.5248
3.6576 | -98.5517
3.6612 | -98.565
3.6648 | -98.555
3.6685 | -98.5342
3.6721 | -98.526
3.6758 | -98.5565
3.6794 | -98.5746
3.6831 | -98.6008
3.6867 | -98.6436
3.6903 | -98.687
3.694 | -98.7267
3.6976 | -98.7572
3.7013 | -98.7681
3.7049 | -98.7569
3.7086 | -98.7253
3.7122 | -98.6808
3.7158 | -98.6308
3.7195 | -98.58
3.7231 | -98.5472
3.7268 | -98.527
3.7304 | -98.5093
3.7341 | -98.5066
3.7377 | -98.5238
3.7413 | -98.5418
3.745 | -98.5745
3.7486 | -98.6145
3.7523 | -98.6659
3.7559 | -98.7356
3.7596 | -98.8027
3.7632 | -98.8637
3.7668 | -98.9056
3.7705 | -98.9175
3.7741 | -98.9043
3.7778 | -98.869
3.7814 | -98.8139
3.7851 | -98.7508
3.7887 | -98.7035
3.7923 | -98.6844
3.796 | -98.6917
3.7996 | -98.7366
3.8033 | -98.8145
3.8069 | -98.892
3.8106 | -98.9596
3.8142 | -98.9993
3.8179 | -98.9933
3.8215 | -98.9496
3.8251 | -98.8703
3.8288 | -98.7796
3.8324 | -98.6896
3.8361 | -98.613
3.8397 | -98.554
3.8434 | -98.5161
3.847 | -98.5036
3.8506 | -98.5139
3.8543 | -98.555
3.8579 | -98.6207
3.8616 | -98.69
3.8652 | -98.7484
3.8689 | -98.7909
3.8725 | -98.8065
3.8761 | -98.7969
3.8798 | -98.7739
3.8834 | -98.7383
3.8871 | -98.7068
3.8907 | -98.6837
3.8944 | -98.6785
3.898 | -98.7039
3.9016 | -98.7681
3.9053 | -98.8602
3.9089 | -98.9738
3.9126 | -99.0952
3.9162 | -99.2083
3.9199 | -99.3005
3.9235 | -99.3699
3.9271 | -99.4125
3.9308 | -99.4247
3.9344 | -99.4148
3.9381 | -99.3841
3.9417 | -99.3246
3.9454 | -99.2492
3.949 | -99.1568
3.9526 | -99.0557
3.9563 | -98.9585
3.9599 | -98.8574
3.9636 | -98.7708
3.9672 | -98.7054
3.9709 | -98.6563
3.9745 | -98.6345
3.9781 | -98.6477
3.9818 | -98.6932
3.9854 | -98.7663
3.9891 | -98.8506
3.9927 | -98.921
3.9964 | -98.9581
4 | -98.9512
4.0036 | -98.917
4.0073 | -98.8794
4.0109 | -98.8505
4.0146 | -98.8505
4.0182 | -98.89
4.0219 | -98.9701
4.0255 | -99.0589
4.0291 | -99.148
4.0328 | -99.2146
4.0364 | -99.2468
4.0401 | -99.227
4.0437 | -99.1687
4.0474 | -99.0891
4.051 | -99.0107
4.0546 | -98.9556
4.0583 | -98.9345
4.0619 | -98.9497
4.0656 | -98.9867
4.0692 | -99.0333
4.0729 | -99.0937
4.0765 | -99.1552
4.0801 | -99.2025
4.0838 | -99.2345
4.0874 | -99.2418
4.0911 | -99.2172
4.0947 | -99.1679
4.0984 | -99.0953
4.102 | -99.0148
4.1056 | -98.9281
4.1093 | -98.8502
4.1129 | -98.8005
4.1166 | -98.7805
4.1202 | -98.7919
4.1239 | -98.8467
4.1275 | -98.932
4.1311 | -99.0324
4.1348 | -99.131
4.1384 | -99.2204
4.1421 | -99.2914
4.1457 | -99.3262
4.1494 | -99.3455
4.153 | -99.3515
4.1566 | -99.3657
4.1603 | -99.398
4.1639 | -99.4574
4.1676 | -99.5437
4.1712 | -99.6421
4.1749 | -99.7363
4.1785 | -99.8253
4.1821 | -99.8888
4.1858 | -99.9182
4.1894 | -99.9114
4.1931 | -99.8767
4.1967 | -99.8216
4.2004 | -99.7526
4.204 | -99.7033
4.2077 | -99.6615
4.2113 | -99.6399
4.2149 | -99.6291
4.2186 | -99.6345
4.2222 | -99.6577
4.2259 | -99.6947
4.2295 | -99.7436
4.2332 | -99.7975
4.2368 | -99.8436
4.2404 | -99.8709
4.2441 | -99.8719
4.2477 | -99.8404
4.2514 | -99.7723
4.255 | -99.6906
4.2587 | -99.5928
4.2623 | -99.4847
4.2659 | -99.3798
4.2696 | -99.2828
4.2732 | -99.1891
4.2769 | -99.1146
4.2805 | -99.0611
4.2842 | -99.0159
4.2878 | -98.9865
4.2914 | -98.9629
4.2951 | -98.9474
4.2987 | -98.9376
4.3024 | -98.9327
4.306 | -98.9312
4.3097 | -98.9295
4.3133 | -98.9274
4.3169 | -98.9259
4.3206 | -98.9247
4.3242 | -98.928
4.3279 | -98.9537
4.3315 | -99.0038
4.3352 | -99.0683
4.3388 | -99.1484
4.3424 | -99.2318
4.3461 | -99.308
4.3497 | -99.3744
4.3534 | -99.4365
4.357 | -99.4865
4.3607 | -99.5386
4.3643 | -99.5931
4.3679 | -99.6408
4.3716 | -99.6831
4.3752 | -99.7212
4.3789 | -99.754
4.3825 | -99.7773
4.3862 | -99.7985
4.3898 | -99.8222
4.3934 | -99.8392
4.3971 | -99.8355
4.4007 | -99.8257
4.4044 | -99.8133
4.408 | -99.7812
4.4117 | -99.7417
4.4153 | -99.7021
4.4189 | -99.6496
4.4226 | -99.5973
4.4262 | -99.5446
4.4299 | -99.4874
4.4335 | -99.4289
4.4372 | -99.3759
4.4408 | -99.3312
4.4444 | -99.3013
4.4481 | -99.288
4.4517 | -99.3022
4.4554 | -99.3509
4.459 | -99.4279
4.4627 | -99.5379
4.4663 | -99.6768
4.4699 | -99.8277
4.4736 | -99.9732
4.4772 | -100.1135
4.4809 | -100.2306
4.4845 | -100.3103
4.4882 | -100.3591
4.4918 | -100.3756
4.4954 | -100.3584
4.4991 | -100.3254
4.5027 | -100.2817
4.5064 | -100.2302
4.51 | -100.1751
4.5137 | -100.1187
4.5173 | -100.0673
4.5209 | -100.0179
4.5246 | -99.9742
4.5282 | -99.9443
4.5319 | -99.9301
4.5355 | -99.9378
4.5392 | -99.9665
4.5428 | -100.0177
4.5464 | -100.0816
4.5501 | -100.1402
4.5537 | -100.1827
4.5574 | -100.1941
4.561 | -100.1638
4.5647 | -100.1014
4.5683 | -100.0035
4.5719 | -99.893
4.5756 | -99.7893
4.5792 | -99.6996
4.5829 | -99.6409
4.5865 | -99.6278
4.5902 | -99.671
4.5938 | -99.7621
4.5975 | -99.8933
4.6011 | -100.0441
4.6047 | -100.2024
4.6084 | -100.3516
4.612 | -100.4711
4.6157 | -100.5479
4.6193 | -100.5695
4.623 | -100.5529
4.6266 | -100.5049
4.6302 | -100.433
4.6339 | -100.3476
4.6375 | -100.2656
4.6412 | -100.192
4.6448 | -100.1425
4.6485 | -100.1325
4.6521 | -100.1544
4.6557 | -100.1986
4.6594 | -100.2582
4.663 | -100.3319
4.6667 | -100.4054
4.6703 | -100.4773
4.674 | -100.531
4.6776 | -100.5707
4.6812 | -100.5985
4.6849 | -100.6253
4.6885 | -100.6576
4.6922 | -100.6773
4.6958 | -100.6894
4.6995 | -100.7003
4.7031 | -100.7041
4.7067 | -100.6974
4.7104 | -100.6844
4.714 | -100.6509
4.7177 | -100.5976
4.7213 | -100.5406
4.725 | -100.4884
4.7286 | -100.4665
4.7322 | -100.4477
4.7359 | -100.4365
4.7395 | -100.4359
4.7432 | -100.4444
4.7468 | -100.4627
4.7505 | -100.4814
4.7541 | -100.5011
4.7577 | -100.5077
4.7614 | -100.5086
4.765 | -100.5039
4.7687 | -100.495
4.7723 | -100.4778
4.776 | -100.4707
4.7796 | -100.4666
4.7832 | -100.4677
4.7869 | -100.48
4.7905 | -100.5042
4.7942 | -100.5401
4.7978 | -100.5817
4.8015 | -100.6237
4.8051 | -100.6628
4.8087 | -100.6949
4.8124 | -100.7181
4.816 | -100.7224
4.8197 | -100.7047
4.8233 | -100.6623
4.827 | -100.6096
4.8306 | -100.5473
4.8342 | -100.4856
4.8379 | -100.4256
4.8415 | -100.3726
4.8452 | -100.3326
4.8488 | -100.3117
4.8525 | -100.3084
4.8561 | -100.3261
4.8597 | -100.3614
4.8634 | -100.4108
4.867 | -100.4697
4.8707 | -100.5255
4.8743 | -100.5849
4.878 | -100.6406
4.8816 | -100.691
4.8852 | -100.7283
4.8889 | -100.751
4.8925 | -100.7411
4.8962 | -100.7074
4.8998 | -100.6534
4.9035 | -100.5781
4.9071 | -100.4922
4.9107 | -100.4144
4.9144 | -100.3461
4.918 | -100.2954
4.9217 | -100.2594
4.9253 | -100.2398
4.929 | -100.2386
4.9326 | -100.2593
4.9362 | -100.3004
4.9399 | -100.356
4.9435 | -100.4197
4.9472 | -100.4909
4.9508 | -100.5546
4.9545 | -100.6078
4.9581 | -100.6466
4.9617 | -100.6636
4.9654 | -100.663
4.969 | -100.661
4.9727 | -100.6445
4.9763 | -100.6302
4.98 | -100.6185
4.9836 | -100.6254
4.9873 | -100.6452
4.9909 | -100.6841
4.9945 | -100.7379
4.9982 | -100.8042
5.0018 | -100.8687
5.0055 | -100.9253
5.0091 | -100.9709
5.0127 | -100.9944
5.0164 | -101.0061
5.02 | -100.9967
5.0237 | -100.9686
5.0273 | -100.9268
5.031 | -100.8785
5.0346 | -100.8312
5.0383 | -100.7943
5.0419 | -100.7647
5.0455 | -100.7358
5.0492 | -100.7219
5.0528 | -100.723
5.0565 | -100.7405
5.0601 | -100.7668
5.0638 | -100.808
5.0674 | -100.8563
5.071 | -100.901
5.0747 | -100.9395
5.0783 | -100.972
5.082 | -100.9944
5.0856 | -101.007
5.0893 | -101.0153
5.0929 | -101.0111
5.0965 | -101.0031
5.1002 | -100.9942
5.1038 | -100.9827
5.1075 | -100.9705
5.1111 | -100.9475
5.1148 | -100.9274
5.1184 | -100.9126
5.122 | -100.9047
5.1257 | -100.899
5.1293 | -100.8992
5.133 | -100.9099
5.1366 | -100.9288
5.1403 | -100.9481
5.1439 | -100.977
5.1475 | -101.0028
5.1512 | -101.0109
5.1548 | -101.0019
5.1585 | -100.9767
5.1621 | -100.939
5.1658 | -100.8887
5.1694 | -100.8355
5.173 | -100.7863
5.1767 | -100.7435
5.1803 | -100.717
5.184 | -100.7027
5.1876 | -100.7035
5.1913 | -100.713
5.1949 | -100.7281
5.1985 | -100.7526
5.2022 | -100.7708
5.2058 | -100.782
5.2095 | -100.7857
5.2131 | -100.7761
5.2168 | -100.7617
5.2204 | -100.7395
5.224 | -100.7231
5.2277 | -100.7162
5.2313 | -100.7101
5.235 | -100.7117
5.2386 | -100.7211
5.2423 | -100.7393
5.2459 | -100.7664
5.2495 | -100.797
5.2532 | -100.8117
5.2568 | -100.8286
5.2605 | -100.839
5.2641 | -100.8454
5.2678 | -100.8499
5.2714 | -100.8547
5.275 | -100.8656
5.2787 | -100.8778
5.2823 | -100.9048
5.286 | -100.9395
5.2896 | -100.9814
5.2933 | -101.0287
5.2969 | -101.0693
5.3005 | -101.106
5.3042 | -101.129
5.3078 | -101.135
5.3115 | -101.1294
5.3151 | -101.0963
5.3188 | -101.0438
5.3224 | -100.9797
5.326 | -100.8999
5.3297 | -100.8156
5.3333 | -100.7344
5.337 | -100.6627
5.3406 | -100.6088
5.3443 | -100.5734
5.3479 | -100.5548
5.3515 | -100.5644
5.3552 | -100.5933
5.3588 | -100.6383
5.3625 | -100.7012
5.3661 | -100.7741
5.3698 | -100.8539
5.3734 | -100.9328
5.377 | -101.013
5.3807 | -101.0844
5.3843 | -101.1461
5.388 | -101.1962
5.3916 | -101.2328
5.3953 | -101.2542
5.3989 | -101.272
5.4025 | -101.2745
5.4062 | -101.2632
5.4098 | -101.2513
5.4135 | -101.2375
5.4171 | -101.2294
5.4208 | -101.2074
5.4244 | -101.1847
5.4281 | -101.1717
5.4317 | -101.1596
5.4353 | -101.1401
5.439 | -101.1262
5.4426 | -101.1181
5.4463 | -101.1118
5.4499 | -101.1124
5.4536 | -101.1169
5.4572 | -101.1268
5.4608 | -101.1305
5.4645 | -101.1318
5.4681 | -101.1263
5.4718 | -101.1266
5.4754 | -101.1279
5.4791 | -101.1286
5.4827 | -101.1345
5.4863 | -101.1477
5.49 | -101.169
5.4936 | -101.2001
5.4973 | -101.2326
5.5009 | -101.2587
5.5046 | -101.2873
5.5082 | -101.3002
5.5118 | -101.2889
5.5155 | -101.2639
5.5191 | -101.2279
5.5228 | -101.1822
5.5264 | -101.1317
5.5301 | -101.0823
5.5337 | -101.0333
5.5373 | -100.9994
5.541 | -100.9805
5.5446 | -100.9823
5.5483 | -101.0009
5.5519 | -101.0227
5.5556 | -101.0468
5.5592 | -101.0774
5.5628 | -101.1104
5.5665 | -101.1413
5.5701 | -101.1705
5.5738 | -101.2052
5.5774 | -101.2368
5.5811 | -101.2738
5.5847 | -101.316
5.5883 | -101.3531
5.592 | -101.3868
5.5956 | -101.4251
5.5993 | -101.4623
5.6029 | -101.4896
5.6066 | -101.5104
5.6102 | -101.532
5.6138 | -101.5527
5.6175 | -101.5651
5.6211 | -101.5663
5.6248 | -101.567
5.6284 | -101.5635
5.6321 | -101.5567
5.6357 | -101.5463
5.6393 | -101.5282
5.643 | -101.5056
5.6466 | -101.4736
5.6503 | -101.4332
5.6539 | -101.3995
5.6576 | -101.3685
5.6612 | -101.3497
5.6648 | -101.347
5.6685 | -101.3596
5.6721 | -101.3902
5.6758 | -101.4311
5.6794 | -101.4752
5.6831 | -101.53
5.6867 | -101.5935
5.6903 | -101.6577
5.694 | -101.7202
5.6976 | -101.7708
5.7013 | -101.816
5.7049 | -101.8401
5.7086 | -101.8403
5.7122 | -101.8311
5.7158 | -101.8128
5.7195 | -101.7846
5.7231 | -101.7474
5.7268 | -101.704
5.7304 | -101.6654
5.7341 | -101.6316
5.7377 | -101.5979
5.7413 | -101.568
5.745 | -101.5541
5.7486 | -101.5544
5.7523 | -101.558
5.7559 | -101.5677
5.7596 | -101.586
5.7632 | -101.6071
5.7668 | -101.622
5.7705 | -101.6335
5.7741 | -101.6488
5.7778 | -101.6648
5.7814 | -101.6798
5.7851 | -101.7003
5.7887 | -101.725
5.7923 | -101.762
5.796 | -101.8055
5.7996 | -101.8494
5.8033 | -101.893
5.8069 | -101.9445
5.8106 | -101.9994
5.8142 | -102.0573
5.8179 | -102.1143
5.8215 | -102.1612
5.8251 | -102.2005
5.8288 | -102.2332
5.8324 | -102.2492
5.8361 | -102.2477
5.8397 | -102.2391
5.8434 | -102.2151
5.847 | -102.1777
5.8506 | -102.1352
5.8543 | -102.0916
5.8579 | -102.0461
5.8616 | -101.9983
5.8652 | -101.9637
5.8689 | -101.9362
5.8725 | -101.9147
5.8761 | -101.8999
5.8798 | -101.8909
5.8834 | -101.8877
5.8871 | -101.8845
5.8907 | -101.8833
5.8944 | -101.8854
5.898 | -101.8851
5.9016 | -101.885
5.9053 | -101.8768
5.9089 | -101.8632
5.9126 | -101.8493
5.9162 | -101.825
5.9199 | -101.7972
5.9235 | -101.7701
5.9271 | -101.7412
5.9308 | -101.7103
5.9344 | -101.6771
5.9381 | -101.6459
5.9417 | -101.6153
5.9454 | -101.5816
5.949 | -101.5412
5.9526 | -101.5023
5.9563 | -101.463
5.9599 | -101.4276
5.9636 | -101.399
5.9672 | -101.378
5.9709 | -101.3652
5.9745 | -101.3601
5.9781 | -101.3647
5.9818 | -101.3803
5.9854 | -101.4033
5.9891 | -101.4302
5.9927 | -101.4547
5.9964 | -101.4907
6 | -101.5247
6.0036 | -101.5521
6.0073 | -101.5844
6.0109 | -101.6038
6.0146 | -101.6164
6.0182 | -101.6258
6.0219 | -101.6269
6.0255 | -101.6248
6.0291 | -101.6214
6.0328 | -101.6211
6.0364 | -101.6354
6.0401 | -101.6637
6.0437 | -101.6962
6.0474 | -101.7335
6.051 | -101.7862
6.0546 | -101.8458
6.0583 | -101.9051
6.0619 | -101.9598
6.0656 | -102.0109
6.0692 | -102.0546
6.0729 | -102.0891
6.0765 | -102.1187
6.0801 | -102.1441
6.0838 | -102.1661
6.0874 | -102.1813
6.0911 | -102.1935
6.0947 | -102.2063
6.0984 | -102.2127
6.102 | -102.2317
6.1056 | -102.2519
6.1093 | -102.2756
6.1129 | -102.3064
6.1166 | -102.3427
6.1202 | -102.3804
6.1239 | -102.4097
6.1275 | -102.4297
6.1311 | -102.4406
6.1348 | -102.4461
6.1384 | -102.4476
6.1421 | -102.4394
6.1457 | -102.4334
6.1494 | -102.4333
6.153 | -102.4332
6.1566 | -102.4356
6.1603 | -102.4432
6.1639 | -102.4472
6.1676 | -102.443
6.1712 | -102.4378
6.1749 | -102.4267
6.1785 | -102.4044
6.1821 | -102.3818
6.1858 | -102.354
6.1894 | -102.3156
6.1931 | -102.2687
6.1967 | -102.2263
6.2004 | -102.1898
6.204 | -102.1629
6.2077 | -102.1477
6.2113 | -102.1469
6.2149 | -102.1607
6.2186 | -102.1919
6.2222 | -102.2412
6.2259 | -102.2998
6.2295 | -102.3625
6.2332 | -102.4264
6.2368 | -102.4793
6.2404 | -102.5189
6.2441 | -102.5428
6.2477 | -102.5414
6.2514 | -102.5141
6.255 | -102.4704
6.2587 | -102.4148
6.2623 | -102.3521
6.2659 | -102.2809
6.2696 | -102.2068
6.2732 | -102.1423
6.2769 | -102.0888
6.2805 | -102.0493
6.2842 | -102.0203
6.2878 | -102
6.2914 | -101.993
6.2951 | -101.9904
6.2987 | -101.9871
6.3024 | -101.9892
6.306 | -101.9884
6.3097 | -101.9957
6.3133 | -101.9998
6.3169 | -101.998
6.3206 | -101.9998
6.3242 | -102.0069
6.3279 | -102.0224
6.3315 | -102.0489
6.3352 | -102.0851
6.3388 | -102.1325
6.3424 | -102.1893
6.3461 | -102.2504
6.3497 | -102.3136
6.3534 | -102.3821
6.357 | -102.4503
6.3607 | -102.5163
6.3643 | -102.5703
6.3679 | -102.6037
6.3716 | -102.6282
6.3752 | -102.6356
6.3789 | -102.6308
6.3825 | -102.6171
6.3862 | -102.5935
6.3898 | -102.5673
6.3934 | -102.5443
6.3971 | -102.5241
6.4007 | -102.5099
6.4044 | -102.5128
6.408 | -102.5239
6.4117 | -102.5456
6.4153 | -102.5831
6.4189 | -102.6274
6.4226 | -102.6739
6.4262 | -102.7127
6.4299 | -102.75
6.4335 | -102.779
6.4372 | -102.7965
6.4408 | -102.8033
6.4444 | -102.7912
6.4481 | -102.7598
6.4517 | -102.7127
6.4554 | -102.6559
6.459 | -102.5961
6.4627 | -102.5321
6.4663 | -102.4615
6.4699 | -102.3841
6.4736 | -102.314
6.4772 | -102.2434
6.4809 | -102.1793
6.4845 | -102.13
6.4882 | -102.094
6.4918 | -102.0696
6.4954 | -102.0549
6.4991 | -102.0528
6.5027 | -102.0642
6.5064 | -102.0915
6.51 | -102.1285
6.5137 | -102.1715
6.5173 | -102.2246
6.5209 | -102.2913
6.5246 | -102.3656
6.5282 | -102.4347
6.5319 | -102.498
6.5355 | -102.5584
6.5392 | -102.6125
6.5428 | -102.6544
6.5464 | -102.6714
6.5501 | -102.6666
6.5537 | -102.639
6.5574 | -102.5922
6.561 | -102.5353
6.5647 | -102.4738
6.5683 | -102.4133
6.5719 | -102.3552
6.5756 | -102.31
6.5792 | -102.2752
6.5829 | -102.2518
6.5865 | -102.2483
6.5902 | -102.261
6.5938 | -102.2873
6.5975 | -102.3244
6.6011 | -102.3696
6.6047 | -102.4183
6.6084 | -102.4717
6.612 | -102.5231
6.6157 | -102.5627
6.6193 | -102.5904
6.623 | -102.6162
6.6266 | -102.6398
6.6302 | -102.659
6.6339 | -102.6636
6.6375 | -102.6619
6.6412 | -102.6625
6.6448 | -102.6679
6.6485 | -102.6712
6.6521 | -102.6695
6.6557 | -102.6702
6.6594 | -102.6722
6.663 | -102.6738
6.6667 | -102.6698
6.6703 | -102.6712
6.674 | -102.6724
6.6776 | -102.6718
6.6812 | -102.6754
6.6849 | -102.6809
6.6885 | -102.6865
6.6922 | -102.694
6.6958 | -102.7045
6.6995 | -102.7229
6.7031 | -102.7448
6.7067 | -102.7679
6.7104 | -102.7986
6.714 | -102.8383
6.7177 | -102.8817
6.7213 | -102.9198
6.725 | -102.9467
6.7286 | -102.9737
6.7322 | -102.9976
6.7359 | -103.0107
6.7395 | -103.0091
6.7432 | -102.9948
6.7468 | -102.9763
6.7505 | -102.9559
6.7541 | -102.9238
6.7577 | -102.8838
6.7614 | -102.84
6.765 | -102.794
6.7687 | -102.7498
6.7723 | -102.7111
6.776 | -102.6804
6.7796 | -102.6563
6.7832 | -102.6398
6.7869 | -102.6367
6.7905 | -102.636
6.7942 | -102.6339
6.7978 | -102.6328
6.8015 | -102.6324
6.8051 | -102.6314
6.8087 | -102.6304
6.8124 | -102.6271
6.816 | -102.6213
6.8197 | -102.6177
6.8233 | -102.6179
6.827 | -102.606
6.8306 | -102.5905
6.8342 | -102.5741
6.8379 | -102.5551
6.8415 | -102.5389
6.8452 | -102.5209
6.8488 | -102.5014
6.8525 | -102.4871
6.8561 | -102.4785
6.8597 | -102.4754
6.8634 | -102.4768
6.867 | -102.4855
6.8707 | -102.497
6.8743 | -102.51
6.878 | -102.5299
6.8816 | -102.5544
6.8852 | -102.5843
6.8889 | -102.6173
6.8925 | -102.6605
6.8962 | -102.7061
6.8998 | -102.7506
6.9035 | -102.8003
6.9071 | -102.8451
6.9107 | -102.8804
6.9144 | -102.9122
6.918 | -102.9372
6.9217 | -102.9438
6.9253 | -102.9461
6.929 | -102.9457
6.9326 | -102.9306
6.9362 | -102.9065
6.9399 | -102.8807
6.9435 | -102.8484
6.9472 | -102.8119
6.9508 | -102.7756
6.9545 | -102.7417
6.9581 | -102.71
6.9617 | -102.6859
6.9654 | -102.6681
6.969 | -102.6588
6.9727 | -102.6562
6.9763 | -102.6589
6.98 | -102.6677
6.9836 | -102.6751
6.9873 | -102.6923
6.9909 | -102.7147
6.9945 | -102.7418
6.9982 | -102.768
7.0018 | -102.7954
7.0055 | -102.824
7.0091 | -102.8445
7.0127 | -102.8661
7.0164 | -102.8872
7.02 | -102.901
7.0237 | -102.9072
7.0273 | -102.9136
7.031 | -102.9072
7.0346 | -102.8966
7.0383 | -102.8848
7.0419 | -102.8646
7.0455 | -102.8431
7.0492 | -102.8181
7.0528 | -102.7901
7.0565 | -102.7571
7.0601 | -102.7266
7.0638 | -102.7026
7.0674 | -102.6825
7.071 | -102.6683
7.0747 | -102.664
7.0783 | -102.6654
7.082 | -102.6665
7.0856 | -102.6789
7.0893 | -102.6964
7.0929 | -102.7063
7.0965 | -102.7164
7.1002 | -102.7262
7.1038 | -102.7261
7.1075 | -102.7183
7.1111 | -102.7018
7.1148 | -102.6778
7.1184 | -102.6465
7.122 | -102.6066
7.1257 | -102.5663
7.1293 | -102.5315
7.133 | -102.4966
7.1366 | -102.4694
7.1403 | -102.4477
7.1439 | -102.4346
7.1475 | -102.434
7.1512 | -102.4453
7.1548 | -102.4732
7.1585 | -102.5109
7.1621 | -102.5648
7.1658 | -102.6302
7.1694 | -102.704
7.173 | -102.7852
7.1767 | -102.8708
7.1803 | -102.9576
7.184 | -103.0452
7.1876 | -103.1322
7.1913 | -103.2075
7.1949 | -103.2737
7.1985 | -103.3275
7.2022 | -103.3652
7.2058 | -103.3866
7.2095 | -103.4008
7.2131 | -103.3962
7.2168 | -103.3775
7.2204 | -103.3513
7.224 | -103.3048
7.2277 | -103.2508
7.2313 | -103.1934
7.235 | -103.128
7.2386 | -103.0644
7.2423 | -103.0011
7.2459 | -102.9435
7.2495 | -102.8973
7.2532 | -102.8567
7.2568 | -102.8235
7.2605 | -102.7956
7.2641 | -102.7779
7.2678 | -102.7663
7.2714 | -102.7626
7.275 | -102.7641
7.2787 | -102.775
7.2823 | -102.7916
7.286 | -102.8123
7.2896 | -102.8348
7.2933 | -102.8628
7.2969 | -102.8959
7.3005 | -102.9287
7.3042 | -102.9546
7.3078 | -102.9797
7.3115 | -103.0025
7.3151 | -103.0188
7.3188 | -103.0276
7.3224 | -103.0254
7.326 | -103.0196
7.3297 | -103.0142
7.3333 | -103.0019
7.337 | -102.9885
7.3406 | -102.98
7.3443 | -102.9743
7.3479 | -102.9731
7.3515 | -102.9817
7.3552 | -103.0102
7.3588 | -103.0293
7.3625 | -103.0536
7.3661 | -103.0819
7.3698 | -103.1125
7.3734 | -103.1448
7.377 | -103.1807
7.3807 | -103.2179
7.3843 | -103.2616
7.388 | -103.3045
7.3916 | -103.3501
7.3953 | -103.4007
7.3989 | -103.456
7.4025 | -103.5067
7.4062 | -103.555
7.4098 | -103.5999
7.4135 | -103.6499
7.4171 | -103.6879
7.4208 | -103.7166
7.4244 | -103.7401
7.4281 | -103.7566
7.4317 | -103.7655
7.4353 | -103.7596
7.439 | -103.746
7.4426 | -103.7338
7.4463 | -103.7201
7.4499 | -103.7035
7.4536 | -103.6827
7.4572 | -103.6545
7.4608 | -103.6287
7.4645 | -103.6025
7.4681 | -103.5829
7.4718 | -103.5624
7.4754 | -103.5405
7.4791 | -103.5174
7.4827 | -103.4941
7.4863 | -103.4606
7.49 | -103.4197
7.4936 | -103.3824
7.4973 | -103.3412
7.5009 | -103.3029
7.5046 | -103.2615
7.5082 | -103.2156
7.5118 | -103.1727
7.5155 | -103.1314
7.5191 | -103.0938
7.5228 | -103.0662
7.5264 | -103.0421
7.5301 | -103.0269
7.5337 | -103.0171
7.5373 | -103.0161
7.541 | -103.027
7.5446 | -103.045
7.5483 | -103.0689
7.5519 | -103.0971
7.5556 | -103.122
7.5592 | -103.1459
7.5628 | -103.171
7.5665 | -103.1979
7.5701 | -103.2206
7.5738 | -103.2307
7.5774 | -103.2392
7.5811 | -103.2366
7.5847 | -103.2283
7.5883 | -103.2203
7.592 | -103.213
7.5956 | -103.2035
7.5993 | -103.2012
7.6029 | -103.203
7.6066 | -103.1958
7.6102 | -103.1993
7.6138 | -103.2043
7.6175 | -103.216
7.6211 | -103.2389
7.6248 | -103.2621
7.6284 | -103.2883
7.6321 | -103.3129
7.6357 | -103.3356
7.6393 | -103.3547
7.643 | -103.3693
7.6466 | -103.3847
7.6503 | -103.3906
7.6539 | -103.3901
7.6576 | -103.3871
7.6612 | -103.3863
7.6648 | -103.3864
7.6685 | -103.387
7.6721 | -103.388
7.6758 | -103.396
7.6794 | -103.4054
7.6831 | -103.4192
7.6867 | -103.4352
7.6903 | -103.4505
7.694 | -103.4751
7.6976 | -103.5022
7.7013 | -103.5255
7.7049 | -103.5473
7.7086 | -103.5759
7.7122 | -103.6065
7.7158 | -103.6345
7.7195 | -103.6656
7.7231 | -103.6877
7.7268 | -103.7087
7.7304 | -103.7236
7.7341 | -103.7315
7.7377 | -103.7397
7.7413 | -103.744
7.745 | -103.7419
7.7486 | -103.7416
7.7523 | -103.7385
7.7559 | -103.7352
7.7596 | -103.7325
7.7632 | -103.7348
7.7668 | -103.7361
7.7705 | -103.7339
7.7741 | -103.7272
7.7778 | -103.7192
7.7814 | -103.71
7.7851 | -103.6948
7.7887 | -103.6767
7.7923 | -103.6598
7.796 | -103.6423
7.7996 | -103.628
7.8033 | -103.6149
7.8069 | -103.5923
7.8106 | -103.5643
7.8142 | -103.5476
7.8179 | -103.5269
7.8215 | -103.5023
7.8251 | -103.477
7.8288 | -103.4519
7.8324 | -103.4274
7.8361 | -103.4077
7.8397 | -103.3887
7.8434 | -103.3704
7.847 | -103.3538
7.8506 | -103.3425
7.8543 | -103.3394
7.8579 | -103.3423
7.8616 | -103.349
7.8652 | -103.3649
7.8689 | -103.3856
7.8725 | -103.4044
7.8761 | -103.4215
7.8798 | -103.4353
7.8834 | -103.4351
7.8871 | -103.4358
7.8907 | -103.4326
7.8944 | -103.4228
7.898 | -103.4097
7.9016 | -103.3936
7.9053 | -103.3764
7.9089 | -103.358
7.9126 | -103.3446
7.9162 | -103.3339
7.9199 | -103.3243
7.9235 | -103.3286
7.9271 | -103.3455
7.9308 | -103.3631
7.9344 | -103.3872
7.9381 | -103.415
7.9417 | -103.4516
7.9454 | -103.4963
7.949 | -103.5449
7.9526 | -103.5927
7.9563 | -103.6426
7.9599 | -103.6922
7.9636 | -103.7359
7.9672 | -103.772
7.9709 | -103.7951
7.9745 | -103.8102
7.9781 | -103.8148
7.9818 | -103.8167
7.9854 | -103.813
7.9891 | -103.8086
7.9927 | -103.802
7.9964 | -103.7946
8 | -103.7823
8.0036 | -103.7703
8.0073 | -103.7683
8.0109 | -103.7666
8.0146 | -103.7649
8.0182 | -103.7649
8.0219 | -103.7736
8.0255 | -103.7843
8.0291 | -103.8022
8.0328 | -103.8224
8.0364 | -103.8448
8.0401 | -103.8724
8.0437 | -103.8994
8.0474 | -103.9193
8.051 | -103.933
8.0546 | -103.9495
8.0583 | -103.9648
8.0619 | -103.9748
8.0656 | -103.9765
8.0692 | -103.9733
8.0729 | -103.9674
8.0765 | -103.9589
8.0801 | -103.9482
8.0838 | -103.9335
8.0874 | -103.9221
8.0911 | -103.9153
8.0947 | -103.9108
8.0984 | -103.908
8.102 | -103.8999
8.1056 | -103.8956
8.1093 | -103.8921
8.1129 | -103.8891
8.1166 | -103.8874
8.1202 | -103.8884
8.1239 | -103.8925
8.1275 | -103.8942
8.1311 | -103.8932
8.1348 | -103.8902
8.1384 | -103.8867
8.1421 | -103.876
8.1457 | -103.8654
8.1494 | -103.8543
8.153 | -103.8412
8.1566 | -103.8275
8.1603 | -103.8188
8.1639 | -103.8069
8.1676 | -103.7988
8.1712 | -103.7948
8.1749 | -103.7904
8.1785 | -103.7829
8.1821 | -103.7734
8.1858 | -103.7706
8.1894 | -103.77
8.1931 | -103.7679
8.1967 | -103.768
8.2004 | -103.7748
8.204 | -103.7875
8.2076 | -103.7964
8.2113 | -103.8123
8.2149 | -103.8301
8.2186 | -103.8483
8.2222 | -103.8698
8.2259 | -103.8917
8.2295 | -103.9168
8.2332 | -103.9378
8.2368 | -103.9548
8.2404 | -103.9705
8.2441 | -103.982
8.2477 | -103.9878
8.2514 | -103.9907
8.255 | -103.9894
8.2587 | -103.9716
8.2623 | -103.952
8.2659 | -103.9261
8.2696 | -103.894
8.2732 | -103.857
8.2769 | -103.8154
8.2805 | -103.7694
8.2842 | -103.7232
8.2878 | -103.6732
8.2914 | -103.6201
8.2951 | -103.5668
8.2987 | -103.5201
8.3024 | -103.4776
8.306 | -103.4377
8.3097 | -103.406
8.3133 | -103.378
8.3169 | -103.354
8.3206 | -103.3347
8.3242 | -103.3233
8.3279 | -103.3222
8.3315 | -103.3251
8.3352 | -103.3364
8.3388 | -103.3572
8.3424 | -103.3825
8.3461 | -103.4131
8.3497 | -103.4467
8.3534 | -103.4865
8.357 | -103.5251
8.3607 | -103.562
8.3643 | -103.5991
8.3679 | -103.6371
8.3716 | -103.6758
8.3752 | -103.7117
8.3789 | -103.7529
8.3825 | -103.7865
8.3862 | -103.8196
8.3898 | -103.8537
8.3934 | -103.8903
8.3971 | -103.9238
8.4007 | -103.9526
8.4044 | -103.9835
8.408 | -104.0149
8.4117 | -104.0476
8.4153 | -104.0803
8.4189 | -104.1107
8.4226 | -104.138
8.4262 | -104.1641
8.4299 | -104.1897
8.4335 | -104.2135
8.4372 | -104.2411
8.4408 | -104.2659
8.4444 | -104.2845
8.4481 | -104.3021
8.4517 | -104.319
8.4554 | -104.3332
8.459 | -104.3397
8.4627 | -104.3386
8.4663 | -104.3319
8.4699 | -104.3211
8.4736 | -104.3047
8.4772 | -104.2806
8.4809 | -104.2553
8.4845 | -104.2233
8.4882 | -104.1839
8.4918 | -104.1424
8.4954 | -104.1008
8.4991 | -104.0625
8.5027 | -104.0252
8.5064 | -103.989
8.51 | -103.9539
8.5137 | -103.9221
8.5173 | -103.8932
8.5209 | -103.869
8.5246 | -103.8513
8.5282 | -103.8322
8.5319 | -103.8208
8.5355 | -103.8111
8.5392 | -103.8058
8.5428 | -103.8056
8.5464 | -103.8088
8.5501 | -103.8163
8.5537 | -103.8296
8.5574 | -103.8426
8.561 | -103.8553
8.5647 | -103.8676
8.5683 | -103.8768
8.5719 | -103.8864
8.5756 | -103.8886
8.5792 | -103.8887
8.5829 | -103.8918
8.5865 | -103.8974
8.5902 | -103.9088
8.5938 | -103.9263
8.5975 | -103.9456
8.6011 | -103.9686
8.6047 | -103.9964
8.6084 | -104.0226
8.612 | -104.0525
8.6157 | -104.0851
8.6193 | -104.1133
8.623 | -104.1385
8.6266 | -104.1648
8.6302 | -104.1842
8.6339 | -104.2003
8.6375 | -104.2127
8.6412 | -104.2192
8.6448 | -104.2259
8.6485 | -104.2305
8.6521 | -104.2321
8.6557 | -104.2351
8.6594 | -104.2348
8.663 | -104.2396
8.6667 | -104.2444
8.6703 | -104.2539
8.674 | -104.2678
8.6776 | -104.2811
8.6812 | -104.298
8.6849 | -104.3248
8.6885 | -104.3547
8.6922 | -104.3907
8.6958 | -104.4349
8.6995 | -104.4772
8.7031 | -104.5188
8.7067 | -104.55
8.7104 | -104.5715
8.714 | -104.5921
8.7177 | -104.6047
8.7213 | -104.6093
8.725 | -104.6107
8.7286 | -104.6033
8.7322 | -104.5939
8.7359 | -104.5873
8.7395 | -104.574
8.7432 | -104.5563
8.7468 | -104.5396
8.7505 | -104.5209
8.7541 | -104.4943
8.7577 | -104.4747
8.7614 | -104.4571
8.765 | -104.4425
8.7687 | -104.4295
8.7723 | -104.4148
8.776 | -104.405
8.7796 | -104.3936
8.7832 | -104.3798
8.7869 | -104.3658
8.7905 | -104.3538
8.7942 | -104.3428
8.7978 | -104.3276
8.8015 | -104.3142
8.8051 | -104.3108
8.8087 | -104.3097
8.8124 | -104.3031
8.816 | -104.2994
8.8197 | -104.2961
8.8233 | -104.2983
8.827 | -104.3033
8.8306 | -104.3116
8.8342 | -104.3223
8.8379 | -104.3326
8.8415 | -104.3365
8.8452 | -104.3392
8.8488 | -104.344
8.8525 | -104.3468
8.8561 | -104.3531
8.8597 | -104.3577
8.8634 | -104.3597
8.867 | -104.361
8.8707 | -104.3641
8.8743 | -104.3675
8.878 | -104.3702
8.8816 | -104.3764
8.8852 | -104.3806
8.8889 | -104.3897
8.8925 | -104.3971
8.8962 | -104.4034
8.8998 | -104.4053
8.9035 | -104.4046
8.9071 | -104.4061
8.9107 | -104.4065
8.9144 | -104.4016
8.918 | -104.394
8.9217 | -104.3887
8.9253 | -104.3863
8.929 | -104.3801
8.9326 | -104.3733
8.9362 | -104.3651
8.9399 | -104.3625
8.9435 | -104.3679
8.9472 | -104.3766
8.9508 | -104.3853
8.9545 | -104.3928
8.9581 | -104.409
8.9617 | -104.4292
8.9654 | -104.4551
8.969 | -104.4833
8.9727 | -104.5181
8.9763 | -104.5566
8.98 | -104.5946
8.9836 | -104.6319
8.9872 | -104.6717
8.9909 | -104.7152
8.9945 | -104.7603
8.9982 | -104.8105
9.0018 | -104.8587
9.0055 | -104.8987
9.0091 | -104.9369
9.0128 | -104.9699
9.0164 | -104.9959
9.02 | -105.0104
9.0237 | -105.0204
9.0273 | -105.0282
9.031 | -105.0319
9.0346 | -105.029
9.0383 | -105.021
9.0419 | -105.0107
9.0455 | -104.9962
9.0492 | -104.9752
9.0528 | -104.9506
9.0565 | -104.923
9.0601 | -104.8961
9.0638 | -104.8655
9.0674 | -104.8283
9.071 | -104.7914
9.0747 | -104.7527
9.0783 | -104.716
9.082 | -104.6793
9.0856 | -104.6441
9.0893 | -104.6054
9.0929 | -104.5697
9.0965 | -104.5392
9.1002 | -104.5131
9.1038 | -104.4917
9.1075 | -104.4706
9.1111 | -104.4508
9.1148 | -104.4383
9.1184 | -104.4345
9.122 | -104.435
9.1257 | -104.4392
9.1293 | -104.4452
9.133 | -104.4529
9.1366 | -104.4636
9.1403 | -104.4818
9.1439 | -104.5041
9.1475 | -104.5301
9.1512 | -104.558
9.1548 | -104.5868
9.1585 | -104.6148
9.1621 | -104.6443
9.1658 | -104.6733
9.1694 | -104.6984
9.173 | -104.7225
9.1767 | -104.7448
9.1803 | -104.764
9.184 | -104.7789
9.1876 | -104.7916
9.1913 | -104.8057
9.1949 | -104.8143
9.1985 | -104.8196
9.2022 | -104.822
9.2058 | -104.8138
9.2095 | -104.8052
9.2131 | -104.7931
9.2168 | -104.7754
9.2204 | -104.7511
9.224 | -104.7282
9.2277 | -104.6985
9.2313 | -104.6659
9.235 | -104.6284
9.2386 | -104.5908
9.2423 | -104.558
9.2459 | -104.5251
9.2495 | -104.4961
9.2532 | -104.4717
9.2568 | -104.4515
9.2605 | -104.4361
9.2641 | -104.4286
9.2678 | -104.4284
9.2714 | -104.4349
9.275 | -104.442
9.2787 | -104.4486
9.2823 | -104.4565
9.286 | -104.4682
9.2896 | -104.4823
9.2933 | -104.4987
9.2969 | -104.5143
9.3005 | -104.5312
9.3042 | -104.5465
9.3078 | -104.5563
9.3115 | -104.5637
9.3151 | -104.5734
9.3188 | -104.5803
9.3224 | -104.5799
9.326 | -104.5731
9.3297 | -104.5649
9.3333 | -104.5582
9.337 | -104.5482
9.3406 | -104.538
9.3443 | -104.5215
9.3479 | -104.5056
9.3515 | -104.4943
9.3552 | -104.4829
9.3588 | -104.4714
9.3625 | -104.4622
9.3661 | -104.4559
9.3698 | -104.4527
9.3734 | -104.4503
9.377 | -104.4481
9.3807 | -104.4487
9.3843 | -104.4434
9.388 | -104.4361
9.3916 | -104.4295
9.3953 | -104.4234
9.3989 | -104.4181
9.4025 | -104.4144
9.4062 | -104.4129
9.4098 | -104.4133
9.4135 | -104.4168
9.4171 | -104.4254
9.4208 | -104.4349
9.4244 | -104.4491
9.4281 | -104.4629
9.4317 | -104.477
9.4353 | -104.4978
9.439 | -104.5215
9.4426 | -104.5427
9.4463 | -104.5615
9.4499 | -104.5719
9.4536 | -104.5809
9.4572 | -104.5716
9.4608 | -104.5694
9.4645 | -104.5599
9.4681 | -104.5493
9.4718 | -104.5312
9.4754 | -104.5091
9.4791 | -104.4816
9.4827 | -104.4522
9.4863 | -104.4205
9.49 | -104.3911
9.4936 | -104.366
9.4973 | -104.3407
9.5009 | -104.3149
9.5046 | -104.2945
9.5082 | -104.2836
9.5118 | -104.2787
9.5155 | -104.277
9.5191 | -104.2825
9.5228 | -104.2961
9.5264 | -104.3172
9.5301 | -104.3458
9.5337 | -104.3813
9.5373 | -104.4207
9.541 | -104.4622
9.5446 | -104.5032
9.5483 | -104.5434
9.5519 | -104.5831
9.5556 | -104.6178
9.5592 | -104.6486
9.5628 | -104.6718
9.5665 | -104.6857
9.5701 | -104.6964
9.5738 | -104.6989
9.5774 | -104.6952
9.5811 | -104.6854
9.5847 | -104.6713
9.5883 | -104.6562
9.592 | -104.6366
9.5956 | -104.6146
9.5993 | -104.5913
9.6029 | -104.5718
9.6066 | -104.5563
9.6102 | -104.5415
9.6138 | -104.5296
9.6175 | -104.5212
9.6211 | -104.5195
9.6248 | -104.5232
9.6284 | -104.5295
9.6321 | -104.5365
9.6357 | -104.5453
9.6393 | -104.5596
9.643 | -104.5773
9.6466 | -104.5956
9.6503 | -104.6119
9.6539 | -104.6307
9.6576 | -104.6513
9.6612 | -104.6721
9.6648 | -104.6904
9.6685 | -104.7053
9.6721 | -104.7139
9.6758 | -104.722
9.6794 | -104.7282
9.6831 | -104.7331
9.6867 | -104.7405
9.6903 | -104.7469
9.694 | -104.7496
9.6976 | -104.7514
9.7013 | -104.7514
9.7049 | -104.7506
9.7086 | -104.7535
9.7122 | -104.7528
9.7158 | -104.749
9.7195 | -104.747
9.7231 | -104.7487
9.7268 | -104.7509
9.7304 | -104.7506
9.7341 | -104.7499
9.7377 | -104.7461
9.7413 | -104.7459
9.745 | -104.7449
9.7486 | -104.7439
9.7523 | -104.744
9.7559 | -104.746
9.7596 | -104.7476
9.7632 | -104.7481
9.7668 | -104.747
9.7705 | -104.7484
9.7741 | -104.7514
9.7778 | -104.7543
9.7814 | -104.7559
9.7851 | -104.7606
9.7887 | -104.7637
9.7924 | -104.7707
9.796 | -104.7789
9.7996 | -104.7902
9.8033 | -104.7982
9.8069 | -104.8071
9.8106 | -104.8178
9.8142 | -104.8313
9.8179 | -104.8395
9.8215 | -104.8474
9.8251 | -104.8537
9.8288 | -104.8559
9.8324 | -104.8572
9.8361 | -104.8548
9.8397 | -104.8505
9.8434 | -104.8461
9.847 | -104.8426
9.8506 | -104.8379
9.8543 | -104.8325
9.8579 | -104.8272
9.8616 | -104.821
9.8652 | -104.8155
9.8689 | -104.8102
9.8725 | -104.8115
9.8761 | -104.8139
9.8798 | -104.8216
9.8834 | -104.8307
9.8871 | -104.8391
9.8907 | -104.848
9.8944 | -104.8555
9.898 | -104.8611
9.9016 | -104.8612
9.9053 | -104.8596
9.9089 | -104.8553
9.9126 | -104.8449
9.9162 | -104.831
9.9199 | -104.8145
9.9235 | -104.7938
9.9271 | -104.7693
9.9308 | -104.7443
9.9344 | -104.7146
9.9381 | -104.6818
9.9417 | -104.6441
9.9454 | -104.6088
9.949 | -104.5775
9.9526 | -104.5424
9.9563 | -104.5077
9.9599 | -104.4765
9.9636 | -104.4471
9.9672 | -104.4177
9.9709 | -104.3946
9.9745 | -104.3773
9.9781 | -104.3608
9.9818 | -104.3456
9.9854 | -104.3478
9.9891 | -104.3411
9.9927 | -104.3414
9.9964 | -104.3467
10 | -104.3599
10.0036 | -104.376
10.0073 | -104.3975
10.0109 | -104.4225
10.0146 | -104.4513
10.0182 | -104.4777
10.0219 | -104.5064
10.0255 | -104.5401
10.0291 | -104.5749
10.0328 | -104.6124
10.0364 | -104.6492
10.0401 | -104.688
10.0437 | -104.728
10.0474 | -104.7659
10.051 | -104.7994
10.0546 | -104.8289
10.0583 | -104.854
10.0619 | -104.8753
10.0656 | -104.8944
10.0692 | -104.9103
10.0729 | -104.9204
10.0765 | -104.924
10.0801 | -104.9263
10.0838 | -104.9289
10.0874 | -104.9329
10.0911 | -104.9304
10.0947 | -104.922
10.0984 | -104.9151
10.102 | -104.9042
10.1056 | -104.8935
10.1093 | -104.884
10.1129 | -104.8709
10.1166 | -104.8568
10.1202 | -104.8447
10.1239 | -104.8385
10.1275 | -104.8344
10.1311 | -104.8291
10.1348 | -104.8243
10.1384 | -104.8195
10.1421 | -104.813
10.1457 | -104.8087
10.1494 | -104.8035
10.153 | -104.8055
10.1566 | -104.81
10.1603 | -104.8129
10.1639 | -104.8167
10.1676 | -104.8296
10.1712 | -104.845
10.1749 | -104.8545
10.1785 | -104.8671
10.1821 | -104.8828
10.1858 | -104.8994
10.1894 | -104.9137
10.1931 | -104.9308
10.1967 | -104.9515
10.2004 | -104.9745
10.204 | -104.9958
10.2076 | -105.0173
10.2113 | -105.0352
10.2149 | -105.0505
10.2186 | -105.0637
10.2222 | -105.0748
10.2259 | -105.0857
10.2295 | -105.0959
10.2332 | -105.0985
10.2368 | -105.1
10.2404 | -105.0989
10.2441 | -105.0924
10.2477 | -105.0832
10.2514 | -105.0737
10.255 | -105.0629
10.2587 | -105.0504
10.2623 | -105.0415
10.2659 | -105.03
10.2696 | -105.0161
10.2732 | -105.0059
10.2769 | -105.0012
10.2805 | -104.9996
10.2842 | -105.0015
10.2878 | -105.0036
10.2914 | -105.0066
10.2951 | -105.0124
10.2987 | -105.0227
10.3024 | -105.0337
10.306 | -105.0441
10.3097 | -105.0577
10.3133 | -105.072
10.3169 | -105.0812
10.3206 | -105.0913
10.3242 | -105.108
10.3279 | -105.1219
10.3315 | -105.1331
10.3352 | -105.1437
10.3388 | -105.1529
10.3424 | -105.1595
10.3461 | -105.1656
10.3497 | -105.1662
10.3534 | -105.1671
10.357 | -105.1642
10.3607 | -105.1604
10.3643 | -105.1549
10.3679 | -105.146
10.3716 | -105.1341
10.3752 | -105.1192
10.3789 | -105.0972
10.3825 | -105.0787
10.3862 | -105.0532
10.3898 | -105.0304
10.3934 | -105.0066
10.3971 | -104.9796
10.4007 | -104.954
10.4044 | -104.9254
10.408 | -104.8991
10.4117 | -104.8728
10.4153 | -104.845
10.4189 | -104.8161
10.4226 | -104.7833
10.4262 | -104.7467
10.4299 | -104.7097
10.4335 | -104.6764
10.4372 | -104.6473
10.4408 | -104.6231
10.4444 | -104.6003
10.4481 | -104.5766
10.4517 | -104.5605
10.4554 | -104.5504
10.459 | -104.547
10.4627 | -104.5508
10.4663 | -104.5628
10.4699 | -104.5805
10.4736 | -104.5993
10.4772 | -104.6229
10.4809 | -104.6529
10.4845 | -104.6851
10.4882 | -104.7181
10.4918 | -104.7515
10.4954 | -104.7801
10.4991 | -104.8045
10.5027 | -104.8236
10.5064 | -104.8377
10.51 | -104.8231
10.5137 | -104.8249
10.5173 | -104.8224
10.5209 | -104.8203
10.5246 | -104.8163
10.5282 | -104.8062
10.5319 | -104.796
10.5355 | -104.7883
10.5392 | -104.7832
10.5428 | -104.7712
10.5464 | -104.7604
10.5501 | -104.7525
10.5537 | -104.743
10.5574 | -104.7342
10.561 | -104.723
10.5647 | -104.7148
10.5683 | -104.712
10.5719 | -104.7135
10.5756 | -104.7157
10.5792 | -104.7223
10.5829 | -104.7293
10.5865 | -104.7326
10.5902 | -104.7389
10.5938 | -104.746
10.5975 | -104.7562
10.6011 | -104.7723
10.6047 | -104.787
10.6084 | -104.8072
10.612 | -104.8292
10.6157 | -104.8524
10.6193 | -104.8762
10.623 | -104.9013
10.6266 | -104.9313
10.6302 | -104.9675
10.6339 | -105.0025
10.6375 | -105.0395
10.6412 | -105.078
10.6448 | -105.1152
10.6485 | -105.1534
10.6521 | -105.1926
10.6557 | -105.23
10.6594 | -105.2669
10.663 | -105.3037
10.6667 | -105.3371
10.6703 | -105.3649
10.674 | -105.3888
10.6776 | -105.4074
10.6812 | -105.4193
10.6849 | -105.4324
10.6885 | -105.4415
10.6922 | -105.438
10.6958 | -105.4325
10.6995 | -105.4294
10.7031 | -105.4191
10.7067 | -105.4039
10.7104 | -105.3883
10.714 | -105.3707
10.7177 | -105.349
10.7213 | -105.3226
10.725 | -105.2955
10.7286 | -105.2683
10.7322 | -105.2426
10.7359 | -105.215
10.7395 | -105.1901
10.7432 | -105.1697
10.7468 | -105.147
10.7505 | -105.1278
10.7541 | -105.1103
10.7577 | -105.0929
10.7614 | -105.0789
10.765 | -105.0673
10.7687 | -105.0578
10.7723 | -105.0534
10.776 | -105.0505
10.7796 | -105.0505
10.7832 | -105.0551
10.7869 | -105.0609
10.7905 | -105.0686
10.7942 | -105.0758
10.7978 | -105.0857
10.8015 | -105.0967
10.8051 | -105.1088
10.8087 | -105.1198
10.8124 | -105.1305
10.816 | -105.1407
10.8197 | -105.1493
10.8233 | -105.1553
10.827 | -105.1633
10.8306 | -105.1682
10.8342 | -105.1722
10.8379 | -105.1768
10.8415 | -105.1775
10.8452 | -105.1777
10.8488 | -105.1776
10.8525 | -105.1741
10.8561 | -105.1704
10.8597 | -105.1682
10.8634 | -105.1599
10.867 | -105.1459
10.8707 | -105.1298
10.8743 | -105.1185
10.878 | -105.1075
10.8816 | -105.0953
10.8852 | -105.0832
10.8889 | -105.0644
10.8925 | -105.0456
10.8962 | -105.0267
10.8998 | -105.007
10.9035 | -104.9888
10.9071 | -104.9704
10.9107 | -104.951
10.9144 | -104.9316
10.918 | -104.9102
10.9217 | -104.8902
10.9253 | -104.8705
10.929 | -104.8491
10.9326 | -104.8328
10.9362 | -104.8188
10.9399 | -104.803
10.9435 | -104.7876
10.9472 | -104.7778
10.9508 | -104.7706
10.9545 | -104.7655
10.9581 | -104.7608
10.9617 | -104.7587
10.9654 | -104.76
10.969 | -104.7633
10.9727 | -104.7672
10.9763 | -104.772
10.98 | -104.7743
10.9836 | -104.7755
10.9872 | -104.781
10.9909 | -104.785
10.9945 | -104.792
10.9982 | -104.8
11.0018 | -104.8043
11.0055 | -104.8098
11.0091 | -104.8164
11.0128 | -104.8194
11.0164 | -104.8223
11.02 | -104.8267
11.0237 | -104.8318
11.0273 | -104.8401
11.031 | -104.8446
11.0346 | -104.8542
11.0383 | -104.8557
11.0419 | -104.8569
11.0455 | -104.8575
11.0492 | -104.8641
11.0528 | -104.8739
11.0565 | -104.8839
11.0601 | -104.8911
11.0638 | -104.8968
11.0674 | -104.9029
11.071 | -104.9126
11.0747 | -104.9221
11.0783 | -104.9301
11.082 | -104.932
11.0856 | -104.9333
11.0893 | -104.9367
11.0929 | -104.9408
11.0965 | -104.9436
11.1002 | -104.9442
11.1038 | -104.9404
11.1075 | -104.9384
11.1111 | -104.9382
11.1148 | -104.9355
11.1184 | -104.9335
11.122 | -104.9345
11.1257 | -104.9357
11.1293 | -104.9393
11.133 | -104.9435
11.1366 | -104.9486
11.1403 | -104.9565
11.1439 | -104.9651
11.1475 | -104.9683
11.1512 | -104.9742
11.1548 | -104.9787
11.1585 | -104.9843
11.1621 | -104.992
11.1658 | -104.9985
11.1694 | -105.0036
11.173 | -105.009
11.1767 | -105.0151
11.1803 | -105.0212
11.184 | -105.0268
11.1876 | -105.0293
11.1913 | -105.0302
11.1949 | -105.0275
11.1985 | -105.0264
11.2022 | -105.0292
11.2058 | -105.0282
11.2095 | -105.0251
11.2131 | -105.0203
11.2168 | -105.0109
11.2204 | -105.0004
11.224 | -104.9919
11.2277 | -104.9822
11.2313 | -104.9707
11.235 | -104.9592
11.2386 | -104.9485
11.2423 | -104.9392
11.2459 | -104.9334
11.2495 | -104.926
11.2532 | -104.9153
11.2568 | -104.9045
11.2605 | -104.8965
11.2641 | -104.8902
11.2678 | -104.8881
11.2714 | -104.8842
11.275 | -104.8831
11.2787 | -104.8819
11.2823 | -104.885
11.286 | -104.889
11.2896 | -104.8948
11.2933 | -104.8997
11.2969 | -104.9095
11.3005 | -104.9224
11.3042 | -104.9383
11.3078 | -104.9573
11.3115 | -104.9817
11.3151 | -105.0024
11.3188 | -105.0227
11.3224 | -105.0418
11.326 | -105.0637
11.3297 | -105.0852
11.3333 | -105.1081
11.337 | -105.1275
11.3406 | -105.1469
11.3443 | -105.1687
11.3479 | -105.1864
11.3515 | -105.2032
11.3552 | -105.2204
11.3588 | -105.2367
11.3625 | -105.2484
11.3661 | -105.2554
11.3698 | -105.2634
11.3734 | -105.2712
11.377 | -105.2754
11.3807 | -105.2772
11.3843 | -105.277
11.388 | -105.2734
11.3916 | -105.2694
11.3953 | -105.264
11.3989 | -105.2554
11.4025 | -105.2471
11.4062 | -105.2407
11.4098 | -105.2356
11.4135 | -105.2288
11.4171 | -105.2241
11.4208 | -105.2193
11.4244 | -105.2126
11.4281 | -105.2076
11.4317 | -105.2068
11.4353 | -105.2057
11.439 | -105.2042
11.4426 | -105.2031
11.4463 | -105.1987
11.4499 | -105.1912
11.4536 | -105.1819
11.4572 | -105.1752
11.4608 | -105.1696
11.4645 | -105.1614
11.4681 | -105.1458
11.4718 | -105.1267
11.4754 | -105.1057
11.4791 | -105.0921
11.4827 | -105.0782
11.4863 | -105.0637
11.49 | -105.0472
11.4936 | -105.029
11.4973 | -105.0119
11.5009 | -104.9971
11.5046 | -104.9797
11.5082 | -104.9623
11.5118 | -104.9521
11.5155 | -104.943
11.5191 | -104.9331
11.5228 | -104.9249
11.5264 | -104.916
11.5301 | -104.907
11.5337 | -104.898
11.5373 | -104.8938
11.541 | -104.8906
11.5446 | -104.8917
11.5483 | -104.8929
11.5519 | -104.8959
11.5556 | -104.899
11.5592 | -104.9106
11.5628 | -104.9209
11.5665 | -104.9329
11.5701 | -104.9446
11.5738 | -104.9554
11.5774 | -104.968
11.5811 | -104.9841
11.5847 | -105.0006
11.5883 | -105.0171
11.592 | -105.037
11.5956 | -105.0577
11.5993 | -105.0714
11.6029 | -105.0833
11.6066 | -105.0929
11.6102 | -105.1035
11.6138 | -105.111
11.6175 | -105.1174
11.6211 | -105.1238
11.6248 | -105.128
11.6284 | -105.1253
11.6321 | -105.1252
11.6357 | -105.1293
11.6393 | -105.1282
11.643 | -105.126
11.6466 | -105.1165
11.6503 | -105.107
11.6539 | -105.0984
11.6576 | -105.0923
11.6612 | -105.084
11.6648 | -105.074
11.6685 | -105.0671
11.6721 | -105.0615
11.6758 | -105.0557
11.6794 | -105.0514
11.6831 | -105.0508
11.6867 | -105.0547
11.6903 | -105.059
11.694 | -105.0666
11.6976 | -105.0777
11.7013 | -105.0887
11.7049 | -105.0994
11.7086 | -105.1114
11.7122 | -105.124
11.7158 | -105.1432
11.7195 | -105.1654
11.7231 | -105.1884
11.7268 | -105.2069
11.7304 | -105.2231
11.7341 | -105.2402
11.7377 | -105.2587
11.7413 | -105.2727
11.745 | -105.2856
11.7486 | -105.2978
11.7523 | -105.3063
11.7559 | -105.3116
11.7596 | -105.3129
11.7632 | -105.3112
11.7668 | -105.3071
11.7705 | -105.3048
11.7741 | -105.2978
11.7778 | -105.2888
11.7814 | -105.2764
11.7851 | -105.2627
11.7887 | -105.2464
11.7924 | -105.2269
11.796 | -105.206
11.7996 | -105.1894
11.8033 | -105.173
11.8069 | -105.1588
11.8106 | -105.1492
11.8142 | -105.1388
11.8179 | -105.1309
11.8215 | -105.1244
11.8251 | -105.1237
11.8288 | -105.1221
11.8324 | -105.127
11.8361 | -105.1322
11.8397 | -105.1385
11.8434 | -105.1475
11.847 | -105.1581
11.8506 | -105.1679
11.8543 | -105.1782
11.8579 | -105.1883
11.8616 | -105.2016
11.8652 | -105.2167
11.8689 | -105.2305
11.8725 | -105.2402
11.8761 | -105.2482
11.8798 | -105.2606
11.8834 | -105.2695
11.8871 | -105.2738
11.8907 | -105.2736
11.8944 | -105.2724
11.898 | -105.2745
11.9016 | -105.2745
11.9053 | -105.2677
11.9089 | -105.2605
11.9126 | -105.2561
11.9162 | -105.2498
11.9199 | -105.2407
11.9235 | -105.2303
11.9271 | -105.2193
11.9308 | -105.2107
11.9344 | -105.2031
11.9381 | -105.1969
11.9417 | -105.1928
11.9454 | -105.192
11.949 | -105.1926
11.9526 | -105.1951
11.9563 | -105.2007
11.9599 | -105.2065
11.9636 | -105.2146
11.9672 | -105.2257
11.9709 | -105.2358
11.9745 | -105.248
11.9781 | -105.2604
11.9818 | -105.2752
11.9854 | -105.2945
11.9891 | -105.3161
11.9927 | -105.3379
11.9964 | -105.3536
12 | -105.3731
12.0036 | -105.3961
12.0073 | -105.4179
12.0109 | -105.4392
12.0146 | -105.4604
12.0182 | -105.4791
12.0219 | -105.5012
12.0255 | -105.5251
12.0291 | -105.5489
12.0328 | -105.5678
12.0364 | -105.5803
12.0401 | -105.5924
12.0437 | -105.6043
12.0474 | -105.6126
12.051 | -105.6163
12.0546 | -105.6163
12.0583 | -105.6153
12.0619 | -105.6146
12.0656 | -105.6082
12.0692 | -105.5994
12.0729 | -105.5867
12.0765 | -105.5745
12.0801 | -105.5613
12.0838 | -105.5497
12.0874 | -105.5338
12.0911 | -105.5208
12.0947 | -105.508
12.0984 | -105.4998
12.102 | -105.4905
12.1056 | -105.4764
12.1093 | -105.4605
12.1129 | -105.4462
12.1166 | -105.4335
12.1202 | -105.4203
12.1239 | -105.4072
12.1275 | -105.3958
12.1311 | -105.386
12.1348 | -105.3777
12.1384 | -105.3707
12.1421 | -105.3632
12.1457 | -105.3566
12.1494 | -105.3488
12.153 | -105.3465
12.1566 | -105.3468
12.1603 | -105.3452
12.1639 | -105.3457
12.1676 | -105.3486
12.1712 | -105.3554
12.1749 | -105.361
12.1785 | -105.3673
12.1821 | -105.3771
12.1858 | -105.3869
12.1894 | -105.3969
12.1931 | -105.4091
12.1967 | -105.4207
12.2004 | -105.4324
12.204 | -105.4437
12.2076 | -105.4564
12.2113 | -105.4727
12.2149 | -105.4874
12.2186 | -105.5014
12.2222 | -105.516
12.2259 | -105.5278
12.2295 | -105.5418
12.2332 | -105.5547
12.2368 | -105.5671
12.2404 | -105.582
12.2441 | -105.5938
12.2477 | -105.6055
12.2514 | -105.6163
12.255 | -105.6244
12.2587 | -105.6296
12.2623 | -105.6346
12.2659 | -105.6376
12.2696 | -105.64
12.2732 | -105.64
12.2769 | -105.6387
12.2805 | -105.638
12.2842 | -105.6405
12.2878 | -105.6422
12.2914 | -105.6406
12.2951 | -105.6336
12.2987 | -105.628
12.3024 | -105.6215
12.306 | -105.6142
12.3097 | -105.6014
12.3133 | -105.5884
12.3169 | -105.5758
12.3206 | -105.561
12.3242 | -105.5452
12.3279 | -105.528
12.3315 | -105.5065
12.3352 | -105.4827
12.3388 | -105.4619
12.3424 | -105.4412
12.3461 | -105.4197
12.3497 | -105.3968
12.3534 | -105.3741
12.357 | -105.3514
12.3607 | -105.3309
12.3643 | -105.309
12.3679 | -105.2844
12.3716 | -105.2639
12.3752 | -105.2435
12.3789 | -105.2239
12.3825 | -105.2099
12.3862 | -105.1968
12.3898 | -105.1829
12.3934 | -105.173
12.3971 | -105.1664
12.4007 | -105.1629
12.4044 | -105.1617
12.408 | -105.1623
12.4117 | -105.1643
12.4153 | -105.168
12.4189 | -105.1707
12.4226 | -105.1752
12.4262 | -105.1795
12.4299 | -105.1847
12.4335 | -105.1912
12.4372 | -105.1985
12.4408 | -105.2069
12.4444 | -105.2139
12.4481 | -105.2178
12.4517 | -105.2201
12.4554 | -105.221
12.459 | -105.2238
12.4627 | -105.2264
12.4663 | -105.2256
12.4699 | -105.2228
12.4736 | -105.2201
12.4772 | -105.2162
12.4809 | -105.2133
12.4845 | -105.2078
12.4882 | -105.2035
12.4918 | -105.1977
12.4954 | -105.19
12.4991 | -105.1885
12.5027 | -105.1855
12.5064 | -105.1826
12.51 | -105.1816
12.5137 | -105.1808
12.5173 | -105.1816
12.5209 | -105.1814
12.5246 | -105.1829
12.5282 | -105.1865
12.5319 | -105.1927
12.5355 | -105.1956
12.5392 | -105.196
12.5428 | -105.2021
12.5464 | -105.212
12.5501 | -105.2198
12.5537 | -105.2258
12.5574 | -105.2337
12.561 | -105.2379
12.5647 | -105.2427
12.5683 | -105.2474
12.5719 | -105.2494
12.5756 | -105.2529
12.5792 | -105.2554
12.5829 | -105.258
12.5865 | -105.2586
12.5902 | -105.2622
12.5938 | -105.2649
12.5975 | -105.2682
12.6011 | -105.2686
12.6047 | -105.2681
12.6084 | -105.2688
12.612 | -105.2716
12.6157 | -105.2733
12.6193 | -105.2752
12.623 | -105.2781
12.6266 | -105.2824
12.6302 | -105.2898
12.6339 | -105.2975
12.6375 | -105.3042
12.6412 | -105.308
12.6448 | -105.3114
12.6485 | -105.3227
12.6521 | -105.3358
12.6557 | -105.3489
12.6594 | -105.3612
12.663 | -105.3719
12.6667 | -105.3826
12.6703 | -105.3937
12.674 | -105.4071
12.6776 | -105.4202
12.6812 | -105.4283
12.6849 | -105.439
12.6885 | -105.451
12.6922 | -105.4615
12.6958 | -105.4718
12.6995 | -105.4802
12.7031 | -105.4912
12.7067 | -105.4996
12.7104 | -105.5065
12.714 | -105.5168
12.7177 | -105.5309
12.7213 | -105.5444
12.725 | -105.5532
12.7286 | -105.5627
12.7322 | -105.5698
12.7359 | -105.5779
12.7395 | -105.581
12.7432 | -105.5856
12.7468 | -105.5908
12.7505 | -105.5893
12.7541 | -105.5839
12.7577 | -105.5785
12.7614 | -105.5726
12.765 | -105.5679
12.7687 | -105.5655
12.7723 | -105.558
12.776 | -105.5503
12.7796 | -105.5425
12.7832 | -105.5316
12.7869 | -105.5205
12.7905 | -105.5073
12.7942 | -105.4949
12.7978 | -105.48
12.8015 | -105.463
12.8051 | -105.4466
12.8087 | -105.4288
12.8124 | -105.4143
12.816 | -105.3997
12.8197 | -105.3875
12.8233 | -105.3746
12.827 | -105.3638
12.8306 | -105.3546
12.8342 | -105.3426
12.8379 | -105.3352
12.8415 | -105.3296
12.8452 | -105.3261
12.8488 | -105.3235
12.8525 | -105.323
12.8561 | -105.3228
12.8597 | -105.3227
12.8634 | -105.3263
12.867 | -105.333
12.8707 | -105.3395
12.8743 | -105.3472
12.878 | -105.3561
12.8816 | -105.3659
12.8852 | -105.3725
12.8889 | -105.3729
12.8925 | -105.3771
12.8962 | -105.3817
12.8998 | -105.3827
12.9035 | -105.3795
12.9071 | -105.3764
12.9107 | -105.3764
12.9144 | -105.3752
12.918 | -105.3738
12.9217 | -105.3716
12.9253 | -105.37
12.929 | -105.3669
12.9326 | -105.3618
12.9362 | -105.3609
12.9399 | -105.3585
12.9435 | -105.3504
12.9472 | -105.3428
12.9508 | -105.3348
12.9545 | -105.3267
12.9581 | -105.3205
12.9617 | -105.3131
12.9654 | -105.3076
12.969 | -105.3081
12.9727 | -105.3028
12.9763 | -105.2947
12.98 | -105.284
12.9836 | -105.2755
12.9872 | -105.268
12.9909 | -105.2609
12.9945 | -105.2547
12.9982 | -105.2463
13.0018 | -105.2378
13.0055 | -105.2324
13.0091 | -105.2282
13.0128 | -105.2242
13.0164 | -105.2221
13.02 | -105.2197
13.0237 | -105.2152
13.0273 | -105.2148
13.031 | -105.2163
13.0346 | -105.2169
13.0383 | -105.219
13.0419 | -105.222
13.0455 | -105.2255
13.0492 | -105.2265
13.0528 | -105.2277
13.0565 | -105.227
13.0601 | -105.2283
13.0638 | -105.2314
13.0674 | -105.234
13.071 | -105.2378
13.0747 | -105.2431
13.0783 | -105.25
13.082 | -105.2552
13.0856 | -105.2568
13.0893 | -105.2614
13.0929 | -105.2658
13.0965 | -105.268
13.1002 | -105.2721
13.1038 | -105.2776
13.1075 | -105.2832
13.1111 | -105.2898
13.1148 | -105.2963
13.1184 | -105.3042
13.122 | -105.31
13.1257 | -105.3188
13.1293 | -105.3279
13.133 | -105.3362
13.1366 | -105.3492
13.1403 | -105.355
13.1439 | -105.3599
13.1475 | -105.3663
13.1512 | -105.3755
13.1548 | -105.3819
13.1585 | -105.3845
13.1621 | -105.3887
13.1658 | -105.391
13.1694 | -105.3935
13.173 | -105.3937
13.1767 | -105.3916
13.1803 | -105.386
13.184 | -105.3785
13.1876 | -105.3652
13.1913 | -105.3491
13.1949 | -105.3291
13.1985 | -105.3077
13.2022 | -105.2844
13.2058 | -105.2585
13.2095 | -105.2339
13.2131 | -105.2088
13.2168 | -105.1788
13.2204 | -105.1478
13.224 | -105.1169
13.2277 | -105.0833
13.2313 | -105.051
13.235 | -105.0194
13.2386 | -104.9897
13.2423 | -104.9635
13.2459 | -104.9382
13.2495 | -104.9134
13.2532 | -104.891
13.2568 | -104.8717
13.2605 | -104.8562
13.2641 | -104.8451
13.2678 | -104.833
13.2714 | -104.8236
13.275 | -104.8197
13.2787 | -104.8188
13.2823 | -104.8205
13.286 | -104.8237
13.2896 | -104.8274
13.2933 | -104.8323
13.2969 | -104.8376
13.3005 | -104.8457
13.3042 | -104.8562
13.3078 | -104.8664
13.3115 | -104.8773
13.3151 | -104.8842
13.3188 | -104.8897
13.3224 | -104.8942
13.326 | -104.8976
13.3297 | -104.9033
13.3333 | -104.9064
13.337 | -104.908
13.3406 | -104.9099
13.3443 | -104.911
13.3479 | -104.9099
13.3515 | -104.91
13.3552 | -104.9091
13.3588 | -104.907
13.3625 | -104.9049
13.3661 | -104.9024
13.3698 | -104.8994
13.3734 | -104.8955
13.377 | -104.8915
13.3807 | -104.89
13.3843 | -104.8892
13.388 | -104.8874
13.3916 | -104.8862
13.3953 | -104.8845
13.3989 | -104.8825
13.4025 | -104.8785
13.4062 | -104.8743
13.4098 | -104.8717
13.4135 | -104.8706
13.4171 | -104.8725
13.4208 | -104.876
13.4244 | -104.8788
13.4281 | -104.8832
13.4317 | -104.8873
13.4353 | -104.8914
13.439 | -104.9003
13.4426 | -104.9089
13.4463 | -104.9147
13.4499 | -104.9207
13.4536 | -104.9285
13.4572 | -104.9348
13.4608 | -104.9419
13.4645 | -104.9492
13.4681 | -104.956
13.4718 | -104.9633
13.4754 | -104.9698
13.4791 | -104.9779
13.4827 | -104.9875
13.4863 | -104.9975
13.49 | -105.0096
13.4936 | -105.0234
13.4973 | -105.0344
13.5009 | -105.0474
13.5046 | -105.0638
13.5082 | -105.0855
13.5118 | -105.1088
13.5155 | -105.1319
13.5191 | -105.1532
13.5228 | -105.172
13.5264 | -105.1961
13.5301 | -105.2195
13.5337 | -105.2391
13.5373 | -105.2594
13.541 | -105.2824
13.5446 | -105.3039
13.5483 | -105.3263
13.5519 | -105.3505
13.5556 | -105.3707
13.5592 | -105.3883
13.5628 | -105.4062
13.5665 | -105.4232
13.5701 | -105.4401
13.5738 | -105.4584
13.5774 | -105.4744
13.5811 | -105.4863
13.5847 | -105.4998
13.5883 | -105.5119
13.592 | -105.5241
13.5956 | -105.5394
13.5993 | -105.5538
13.6029 | -105.5638
13.6066 | -105.5696
13.6102 | -105.5739
13.6138 | -105.5788
13.6175 | -105.5833
13.6211 | -105.5844
13.6248 | -105.5857
13.6284 | -105.5861
13.6321 | -105.5831
13.6357 | -105.5786
13.6393 | -105.5747
13.643 | -105.5692
13.6466 | -105.56
13.6503 | -105.5478
13.6539 | -105.5324
13.6576 | -105.5189
13.6612 | -105.4976
13.6648 | -105.4789
13.6685 | -105.4596
13.6721 | -105.4415
13.6758 | -105.4203
13.6794 | -105.3996
13.6831 | -105.3791
13.6867 | -105.357
13.6903 | -105.3352
13.694 | -105.3185
13.6976 | -105.3001
13.7013 | -105.2807
13.7049 | -105.2612
13.7086 | -105.2437
13.7122 | -105.2268
13.7158 | -105.2144
13.7195 | -105.2053
13.7231 | -105.192
13.7268 | -105.1811
13.7304 | -105.1746
13.7341 | -105.1665
13.7377 | -105.1605
13.7413 | -105.1566
13.745 | -105.1525
13.7486 | -105.1492
13.7523 | -105.1499
13.7559 | -105.153
13.7596 | -105.1598
13.7632 | -105.1673
13.7668 | -105.1749
13.7705 | -105.1854
13.7741 | -105.1966
13.7778 | -105.2122
13.7814 | -105.227
13.7851 | -105.2404
13.7887 | -105.2517
13.7924 | -105.2607
13.796 | -105.272
13.7996 | -105.2856
13.8033 | -105.299
13.8069 | -105.3092
13.8106 | -105.3164
13.8142 | -105.321
13.8179 | -105.3235
13.8215 | -105.3219
13.8251 | -105.3194
13.8288 | -105.3153
13.8324 | -105.3078
13.8361 | -105.2984
13.8397 | -105.2896
13.8434 | -105.2789
13.847 | -105.2662
13.8506 | -105.2521
13.8543 | -105.2348
13.8579 | -105.2153
13.8616 | -105.1957
13.8652 | -105.1774
13.8689 | -105.159
13.8725 | -105.1393
13.8761 | -105.1227
13.8798 | -105.108
13.8834 | -105.0966
13.8871 | -105.0883
13.8907 | -105.0801
13.8944 | -105.0742
13.898 | -105.071
13.9016 | -105.0699
13.9053 | -105.0689
13.9089 | -105.0712
13.9126 | -105.0805
13.9162 | -105.0901
13.9199 | -105.1014
13.9235 | -105.1137
13.9271 | -105.1239
13.9308 | -105.1344
13.9344 | -105.1476
13.9381 | -105.1606
13.9417 | -105.1735
13.9454 | -105.1859
13.949 | -105.1976
13.9526 | -105.207
13.9563 | -105.2158
13.9599 | -105.2225
13.9636 | -105.2268
13.9672 | -105.228
13.9709 | -105.2263
13.9745 | -105.2278
13.9781 | -105.2247
13.9818 | -105.2197
13.9854 | -105.2172
13.9891 | -105.2148
13.9927 | -105.2117
13.9964 | -105.2062
14 | -105.1995
14.0036 | -105.1906
14.0073 | -105.1801
14.0109 | -105.17
14.0146 | -105.1613
14.0182 | -105.1524
14.0219 | -105.1429
14.0255 | -105.1291
14.0291 | -105.118
14.0328 | -105.1098
14.0364 | -105.0997
14.0401 | -105.0871
14.0437 | -105.0756
14.0474 | -105.0681
14.051 | -105.0606
14.0546 | -105.0516
14.0583 | -105.046
14.0619 | -105.042
14.0656 | -105.0391
14.0692 | -105.0361
14.0729 | -105.0337
14.0765 | -105.0313
14.0801 | -105.0302
14.0838 | -105.0316
14.0874 | -105.0354
14.0911 | -105.043
14.0947 | -105.0495
14.0984 | -105.0606
14.102 | -105.0743
14.1056 | -105.0913
14.1093 | -105.1079
14.1129 | -105.1214
14.1166 | -105.1375
14.1202 | -105.1538
14.1239 | -105.1737
14.1275 | -105.1983
14.1311 | -105.2203
14.1348 | -105.2399
14.1384 | -105.2604
14.1421 | -105.2813
14.1457 | -105.306
14.1494 | -105.3284
14.153 | -105.3498
14.1566 | -105.371
14.1603 | -105.3924
14.1639 | -105.4153
14.1676 | -105.4407
14.1712 | -105.4637
14.1749 | -105.4839
14.1785 | -105.505
14.1821 | -105.5245
14.1858 | -105.5406
14.1894 | -105.5516
14.1931 | -105.5658
14.1967 | -105.5787
14.2004 | -105.59
14.204 | -105.5982
14.2076 | -105.6068
14.2113 | -105.618
14.2149 | -105.6268
14.2186 | -105.633
14.2222 | -105.6399
14.2259 | -105.6479
14.2295 | -105.6538
14.2332 | -105.6561
14.2368 | -105.6582
14.2404 | -105.6619
14.2441 | -105.6667
14.2477 | -105.6702
14.2514 | -105.671
14.255 | -105.6751
14.2587 | -105.6803
14.2623 | -105.6828
14.2659 | -105.6878
14.2696 | -105.6925
14.2732 | -105.6944
14.2769 | -105.6992
14.2805 | -105.7047
14.2842 | -105.7076
14.2878 | -105.7073
14.2914 | -105.7051
14.2951 | -105.7021
14.2987 | -105.7036
14.3024 | -105.7033
14.306 | -105.7022
14.3097 | -105.7009
14.3133 | -105.6976
14.3169 | -105.693
14.3206 | -105.6875
14.3242 | -105.6839
14.3279 | -105.6814
14.3315 | -105.6729
14.3352 | -105.6613
14.3388 | -105.649
14.3424 | -105.6362
14.3461 | -105.6204
14.3497 | -105.6008
14.3534 | -105.5827
14.357 | -105.564
14.3607 | -105.5434
14.3643 | -105.5208
14.3679 | -105.4994
14.3716 | -105.475
14.3752 | -105.4507
14.3789 | -105.4288
14.3825 | -105.4067
14.3862 | -105.3854
14.3898 | -105.3668
14.3934 | -105.3469
14.3971 | -105.3292
14.4007 | -105.3146
14.4044 | -105.301
14.408 | -105.2902
14.4117 | -105.2822
14.4153 | -105.2745
14.4189 | -105.2668
14.4226 | -105.2619
14.4262 | -105.2596
14.4299 | -105.2602
14.4335 | -105.263
14.4372 | -105.2697
14.4408 | -105.2763
14.4444 | -105.2841
14.4481 | -105.2916
14.4517 | -105.3001
14.4554 | -105.3107
14.459 | -105.3219
14.4627 | -105.3309
14.4663 | -105.3431
14.4699 | -105.3564
14.4736 | -105.3685
14.4772 | -105.382
14.4809 | -105.3962
14.4845 | -105.4068
14.4882 | -105.4171
14.4918 | -105.4289
14.4954 | -105.4389
14.4991 | -105.4476
14.5027 | -105.4522
14.5064 | -105.4524
14.51 | -105.45
14.5137 | -105.446
14.5173 | -105.4389
14.5209 | -105.4327
14.5246 | -105.4249
14.5282 | -105.4167
14.5319 | -105.4043
14.5355 | -105.3869
14.5392 | -105.3705
14.5428 | -105.3537
14.5464 | -105.3373
14.5501 | -105.3216
14.5537 | -105.308
14.5574 | -105.2943
14.561 | -105.2796
14.5647 | -105.2651
14.5683 | -105.2542
14.5719 | -105.2463
14.5756 | -105.2392
14.5792 | -105.2355
14.5829 | -105.2336
14.5865 | -105.2315
14.5902 | -105.2363
14.5938 | -105.2427
14.5975 | -105.2486
14.6011 | -105.2582
14.6047 | -105.2705
14.6084 | -105.2822
14.612 | -105.2975
14.6157 | -105.3127
14.6193 | -105.3317
14.623 | -105.3516
14.6266 | -105.3723
14.6302 | -105.3919
14.6339 | -105.4104
14.6375 | -105.4256
14.6412 | -105.4406
14.6448 | -105.4543
14.6485 | -105.4712
14.6521 | -105.4858
14.6557 | -105.5
14.6594 | -105.51
14.663 | -105.5189
14.6667 | -105.5269
14.6703 | -105.5333
14.674 | -105.5374
14.6776 | -105.5396
14.6812 | -105.5379
14.6849 | -105.5339
14.6885 | -105.5313
14.6922 | -105.5288
14.6958 | -105.5235
14.6995 | -105.5171
14.7031 | -105.5142
14.7067 | -105.5104
14.7104 | -105.5042
14.714 | -105.4908
14.7177 | -105.4819
14.7213 | -105.4735
14.725 | -105.4639
14.7286 | -105.4525
14.7322 | -105.4369
14.7359 | -105.4217
14.7395 | -105.4039
14.7432 | -105.3872
14.7468 | -105.3743
14.7505 | -105.3617
14.7541 | -105.3485
14.7577 | -105.3342
14.7614 | -105.3155
14.765 | -105.2983
14.7687 | -105.2834
14.7723 | -105.268
14.776 | -105.25
14.7796 | -105.2313
14.7832 | -105.2135
14.7869 | -105.1961
14.7905 | -105.181
14.7942 | -105.1676
14.7978 | -105.1544
14.8015 | -105.1385
14.8051 | -105.1219
14.8087 | -105.1051
14.8124 | -105.0892
14.816 | -105.0728
14.8197 | -105.0588
14.8233 | -105.0453
14.827 | -105.033
14.8306 | -105.0169
14.8342 | -104.9989
14.8379 | -104.9831
14.8415 | -104.9677
14.8452 | -104.9489
14.8488 | -104.9335
14.8525 | -104.9188
14.8561 | -104.9023
14.8597 | -104.8808
14.8634 | -104.859
14.867 | -104.836
14.8707 | -104.8137
14.8743 | -104.7897
14.878 | -104.7671
14.8816 | -104.7452
14.8852 | -104.719
14.8889 | -104.6892
14.8925 | -104.6627
14.8962 | -104.6376
14.8998 | -104.6098
14.9035 | -104.5809
14.9071 | -104.5524
14.9107 | -104.5247
14.9144 | -104.4947
14.918 | -104.4646
14.9217 | -104.4396
14.9253 | -104.4172
14.929 | -104.3953
14.9326 | -104.3735
14.9362 | -104.3521
14.9399 | -104.3313
14.9435 | -104.3127
14.9472 | -104.2972
14.9508 | -104.2837
14.9545 | -104.2743
14.9581 | -104.2662
14.9617 | -104.2583
14.9654 | -104.2536
14.969 | -104.2536
14.9727 | -104.254
14.9763 | -104.2566
14.98 | -104.2597
14.9836 | -104.2625
14.9872 | -104.2689
14.9909 | -104.2802
14.9945 | -104.2927
14.9982 | -104.3059
15.0018 | -104.3205
15.0055 | -104.334
15.0091 | -104.3477
15.0128 | -104.3623
15.0164 | -104.3776
15.02 | -104.396
15.0237 | -104.4143
15.0273 | -104.4285
15.031 | -104.4408
15.0346 | -104.4527
15.0383 | -104.4632
15.0419 | -104.474
15.0455 | -104.4864
15.0492 | -104.498
15.0528 | -104.5071
15.0565 | -104.5136
15.0601 | -104.523
15.0638 | -104.5286
15.0674 | -104.5291
15.071 | -104.5304
15.0747 | -104.5339
15.0783 | -104.5358
15.082 | -104.5399
15.0856 | -104.5428
15.0893 | -104.5448
15.0929 | -104.5448
15.0965 | -104.5449
15.1002 | -104.5483
15.1038 | -104.5533
15.1075 | -104.556
15.1111 | -104.5571
15.1148 | -104.5603
15.1184 | -104.564
15.122 | -104.566
15.1257 | -104.5679
15.1293 | -104.5725
15.133 | -104.5794
15.1366 | -104.5864
15.1403 | -104.5923
15.1439 | -104.5999
15.1475 | -104.6088
15.1512 | -104.6186
15.1548 | -104.6317
15.1585 | -104.6475
15.1621 | -104.6631
15.1658 | -104.677
15.1694 | -104.6898
15.173 | -104.705
15.1767 | -104.7218
15.1803 | -104.7374
15.184 | -104.7523
15.1876 | -104.7711
15.1913 | -104.7898
15.1949 | -104.8072
15.1985 | -104.8224
15.2022 | -104.8335
15.2058 | -104.8407
15.2095 | -104.8435
15.2131 | -104.842
15.2168 | -104.8408
15.2204 | -104.8385
15.224 | -104.832
15.2277 | -104.8229
15.2313 | -104.8092
15.235 | -104.7893
15.2386 | -104.752
15.2423 | -104.7279
15.2459 | -104.7034
15.2495 | -104.6777
15.2532 | -104.6474
15.2568 | -104.6134
15.2605 | -104.5792
15.2641 | -104.5455
15.2678 | -104.512
15.2714 | -104.4755
15.275 | -104.4427
15.2787 | -104.41
15.2823 | -104.3786
15.286 | -104.3508
15.2896 | -104.3242
15.2933 | -104.2987
15.2969 | -104.2759
15.3005 | -104.2542
15.3042 | -104.238
15.3078 | -104.2246
15.3115 | -104.2153
15.3151 | -104.2089
15.3188 | -104.2059
15.3224 | -104.2105
15.326 | -104.2189
15.3297 | -104.2292
15.3333 | -104.2412
15.337 | -104.2564
15.3406 | -104.2761
15.3443 | -104.2979
15.3479 | -104.3205
15.3515 | -104.3419
15.3552 | -104.3633
15.3588 | -104.3866
15.3625 | -104.4094
15.3661 | -104.431
15.3698 | -104.4507
15.3734 | -104.4684
15.377 | -104.4813
15.3807 | -104.4915
15.3843 | -104.5007
15.388 | -104.5036
15.3916 | -104.5034
15.3953 | -104.5024
15.3989 | -104.4979
15.4025 | -104.4893
15.4062 | -104.4741
15.4098 | -104.4545
15.4135 | -104.434
15.4171 | -104.413
15.4208 | -104.3891
15.4244 | -104.3613
15.4281 | -104.3319
15.4317 | -104.3011
15.4353 | -104.2652
15.439 | -104.2263
15.4426 | -104.1875
15.4463 | -104.1488
15.4499 | -104.1115
15.4536 | -104.0764
15.4572 | -104.0405
15.4608 | -104.0066
15.4645 | -103.9749
15.4681 | -103.9465
15.4718 | -103.9203
15.4754 | -103.8968
15.4791 | -103.8764
15.4827 | -103.8591
15.4863 | -103.8452
15.49 | -103.8338
15.4936 | -103.8275
15.4973 | -103.8241
15.5009 | -103.8231
15.5046 | -103.8251
15.5082 | -103.8315
15.5118 | -103.84
15.5155 | -103.8505
15.5191 | -103.8646
15.5228 | -103.8805
15.5264 | -103.8992
15.5301 | -103.9212
15.5337 | -103.9439
15.5373 | -103.9682
15.541 | -103.9906
15.5446 | -104.0121
15.5483 | -104.0363
15.5519 | -104.0615
15.5556 | -104.0873
15.5592 | -104.1137
15.5628 | -104.1388
15.5665 | -104.1601
15.5701 | -104.181
15.5738 | -104.2039
15.5774 | -104.2267
15.5811 | -104.2476
15.5847 | -104.268
15.5883 | -104.2879
15.592 | -104.3055
15.5956 | -104.3216
15.5993 | -104.3383
15.6029 | -104.3543
15.6066 | -104.3717
15.6102 | -104.3917
15.6138 | -104.4091
15.6175 | -104.4237
15.6211 | -104.4392
15.6248 | -104.4554
15.6284 | -104.4737
15.6321 | -104.4934
15.6357 | -104.5119
15.6393 | -104.5294
15.643 | -104.5491
15.6466 | -104.5715
15.6503 | -104.5939
15.6539 | -104.6131
15.6576 | -104.6336
15.6612 | -104.6574
15.6648 | -104.683
15.6685 | -104.7081
15.6721 | -104.7307
15.6758 | -104.7517
15.6794 | -104.774
15.6831 | -104.7962
15.6867 | -104.82
15.6903 | -104.8419
15.694 | -104.8621
15.6976 | -104.8827
15.7013 | -104.9047
15.7049 | -104.9248
15.7086 | -104.9412
15.7122 | -104.9587
15.7158 | -104.9756
15.7195 | -104.9913
15.7231 | -105.0064
15.7268 | -105.0198
15.7304 | -105.0323
15.7341 | -105.0473
15.7377 | -105.061
15.7413 | -105.0704
15.745 | -105.0794
15.7486 | -105.0876
15.7523 | -105.095
15.7559 | -105.1024
15.7596 | -105.1104
15.7632 | -105.1215
15.7668 | -105.1278
15.7705 | -105.1351
15.7741 | -105.1392
15.7778 | -105.1461
15.7814 | -105.1521
15.7851 | -105.1571
15.7887 | -105.1597
15.7924 | -105.1647
15.796 | -105.17
15.7996 | -105.1752
15.8033 | -105.1782
15.8069 | -105.1815
15.8106 | -105.1832
15.8142 | -105.1834
15.8179 | -105.1862
15.8215 | -105.1885
15.8251 | -105.1917
15.8288 | -105.1941
15.8324 | -105.1913
15.8361 | -105.1883
15.8397 | -105.1868
15.8434 | -105.1874
15.847 | -105.189
15.8506 | -105.1907
15.8543 | -105.1897
15.8579 | -105.1891
15.8616 | -105.1876
15.8652 | -105.1884
15.8689 | -105.1901
15.8725 | -105.189
15.8761 | -105.1868
15.8798 | -105.1886
15.8834 | -105.1897
15.8871 | -105.1869
15.8907 | -105.1828
15.8944 | -105.1781
15.898 | -105.1758
15.9016 | -105.1688
15.9053 | -105.1601
15.9089 | -105.1525
15.9126 | -105.1432
15.9162 | -105.1307
15.9199 | -105.1194
15.9235 | -105.1103
15.9271 | -105.1017
15.9308 | -105.0918
15.9344 | -105.0831
15.9381 | -105.0722
15.9417 | -105.0569
15.9454 | -105.042
15.949 | -105.0289
15.9526 | -105.019
15.9563 | -105.0082
15.9599 | -104.9954
15.9636 | -104.9843
15.9672 | -104.9779
15.9709 | -104.9728
15.9745 | -104.9651
15.9781 | -104.9556
15.9818 | -104.9475
15.9854 | -104.9409
15.9891 | -104.9333
15.9927 | -104.9268
15.9964 | -104.922
16 | -104.9175
16.0036 | -104.9144
16.0073 | -104.9122
16.0109 | -104.9114
16.0146 | -104.9118
16.0182 | -104.9142
16.0219 | -104.9158
16.0255 | -104.917
16.0291 | -104.918
16.0328 | -104.9202
16.0364 | -104.9226
16.0401 | -104.9282
16.0437 | -104.9358
16.0474 | -104.9414
16.051 | -104.9445
16.0546 | -104.9486
16.0583 | -104.9534
16.0619 | -104.9596
16.0656 | -104.9649
16.0692 | -104.9666
16.0729 | -104.9673
16.0765 | -104.9686
16.0801 | -104.9696
16.0838 | -104.9703
16.0874 | -104.9696
16.0911 | -104.9669
16.0947 | -104.9637
16.0984 | -104.9607
16.102 | -104.9567
16.1056 | -104.9529
16.1093 | -104.9493
16.1129 | -104.9445
16.1166 | -104.9385
16.1202 | -104.9305
16.1239 | -104.9222
16.1275 | -104.9134
16.1311 | -104.9071
16.1348 | -104.8981
16.1384 | -104.8836
16.1421 | -104.8705
16.1457 | -104.8597
16.1494 | -104.8487
16.153 | -104.8372
16.1566 | -104.825
16.1603 | -104.8134
16.1639 | -104.8023
16.1676 | -104.793
16.1712 | -104.7884
16.1749 | -104.7831
16.1785 | -104.7765
16.1821 | -104.7711
16.1858 | -104.7675
16.1894 | -104.7663
16.1931 | -104.7677
16.1967 | -104.7701
16.2004 | -104.7759
16.204 | -104.784
16.2077 | -104.7962
16.2113 | -104.8112
16.2149 | -104.8278
16.2186 | -104.8459
16.2222 | -104.8655
16.2259 | -104.8882
16.2295 | -104.9132
16.2332 | -104.938
16.2368 | -104.9606
16.2404 | -104.9847
16.2441 | -105.0102
16.2477 | -105.0382
16.2514 | -105.0685
16.255 | -105.0965
16.2587 | -105.1233
16.2623 | -105.149
16.2659 | -105.1769
16.2696 | -105.201
16.2732 | -105.2233
16.2769 | -105.246
16.2805 | -105.2677
16.2842 | -105.2844
16.2878 | -105.2987
16.2914 | -105.3006
16.2951 | -105.3149
16.2987 | -105.3239
16.3024 | -105.3332
16.306 | -105.3399
16.3097 | -105.3435
16.3133 | -105.3455
16.3169 | -105.3478
16.3206 | -105.3502
16.3242 | -105.3508
16.3279 | -105.3467
16.3315 | -105.3413
16.3352 | -105.3378
16.3388 | -105.3353
16.3424 | -105.3318
16.3461 | -105.3287
16.3497 | -105.325
16.3534 | -105.3237
16.357 | -105.3222
16.3607 | -105.3212
16.3643 | -105.3221
16.3679 | -105.3253
16.3716 | -105.3322
16.3752 | -105.3382
16.3789 | -105.3435
16.3825 | -105.3489
16.3862 | -105.3541
16.3898 | -105.3596
16.3934 | -105.3687
16.3971 | -105.3803
16.4007 | -105.3919
16.4044 | -105.4014
16.408 | -105.4111
16.4117 | -105.4231
16.4153 | -105.4345
16.4189 | -105.4458
16.4226 | -105.4604
16.4262 | -105.4725
16.4299 | -105.4843
16.4335 | -105.4958
16.4372 | -105.508
16.4408 | -105.5192
16.4444 | -105.5299
16.4481 | -105.5397
16.4517 | -105.5495
16.4554 | -105.5575
16.459 | -105.5635
16.4627 | -105.5673
16.4663 | -105.57
16.4699 | -105.5743
16.4736 | -105.5742
16.4772 | -105.5706
16.4809 | -105.5643
16.4845 | -105.5579
16.4882 | -105.5518
16.4918 | -105.546
16.4954 | -105.5429
16.4991 | -105.5365
16.5027 | -105.5307
16.5064 | -105.5255
16.51 | -105.5196
16.5137 | -105.5157
16.5173 | -105.5137
16.5209 | -105.5119
16.5246 | -105.5111
16.5282 | -105.5085
16.5319 | -105.5062
16.5355 | -105.5061
16.5392 | -105.5079
16.5428 | -105.5122
16.5464 | -105.5187
16.5501 | -105.5245
16.5537 | -105.5291
16.5574 | -105.5352
16.561 | -105.5439
16.5647 | -105.5517
16.5683 | -105.5579
16.5719 | -105.5657
16.5756 | -105.5746
16.5792 | -105.5844
16.5829 | -105.5942
16.5865 | -105.6082
16.5902 | -105.6188
16.5938 | -105.6313
16.5974 | -105.643
16.6011 | -105.6517
16.6047 | -105.6569
16.6084 | -105.6616
16.612 | -105.6684
16.6157 | -105.6744
16.6193 | -105.6804
16.623 | -105.6866
16.6266 | -105.6914
16.6302 | -105.6924
16.6339 | -105.6919
16.6375 | -105.6912
16.6412 | -105.6913
16.6448 | -105.6901
16.6485 | -105.6844
16.6521 | -105.674
16.6557 | -105.6639
16.6594 | -105.6519
16.663 | -105.6421
16.6667 | -105.6301
16.6703 | -105.6149
16.674 | -105.6006
16.6776 | -105.587
16.6812 | -105.5718
16.6849 | -105.5566
16.6885 | -105.5396
16.6922 | -105.5208
16.6958 | -105.5053
16.6995 | -105.4917
16.7031 | -105.4781
16.7067 | -105.4624
16.7104 | -105.4471
16.714 | -105.4349
16.7177 | -105.4252
16.7213 | -105.4134
16.725 | -105.4022
16.7286 | -105.3912
16.7322 | -105.3795
16.7359 | -105.3701
16.7395 | -105.3597
16.7432 | -105.3501
16.7468 | -105.343
16.7505 | -105.3368
16.7541 | -105.3306
16.7577 | -105.3266
16.7614 | -105.3225
16.765 | -105.3176
16.7687 | -105.3159
16.7723 | -105.3176
16.776 | -105.3186
16.7796 | -105.3169
16.7832 | -105.3147
16.7869 | -105.3178
16.7905 | -105.3249
16.7942 | -105.3297
16.7978 | -105.3339
16.8015 | -105.3413
16.8051 | -105.3476
16.8087 | -105.3535
16.8124 | -105.3591
16.816 | -105.3684
16.8197 | -105.3746
16.8233 | -105.3828
16.827 | -105.3901
16.8306 | -105.3982
16.8342 | -105.4047
16.8379 | -105.4092
16.8415 | -105.415
16.8452 | -105.4225
16.8488 | -105.4289
16.8525 | -105.4311
16.8561 | -105.4318
16.8597 | -105.4329
16.8634 | -105.4335
16.867 | -105.4321
16.8707 | -105.4323
16.8743 | -105.4303
16.878 | -105.4268
16.8816 | -105.4209
16.8852 | -105.415
16.8889 | -105.4083
16.8925 | -105.4017
16.8962 | -105.3941
16.8998 | -105.3847
16.9035 | -105.3747
16.9071 | -105.3636
16.9107 | -105.3539
16.9144 | -105.3448
16.918 | -105.335
16.9217 | -105.3252
16.9253 | -105.3147
16.929 | -105.3056
16.9326 | -105.2968
16.9362 | -105.2888
16.9399 | -105.2796
16.9435 | -105.269
16.9472 | -105.2627
16.9508 | -105.2586
16.9545 | -105.251
16.9581 | -105.2444
16.9617 | -105.2389
16.9654 | -105.233
16.969 | -105.2281
16.9727 | -105.2249
16.9763 | -105.2231
16.98 | -105.2213
16.9836 | -105.2172
16.9872 | -105.214
16.9909 | -105.2129
16.9945 | -105.2155
16.9982 | -105.2167
17.0018 | -105.2194
17.0055 | -105.2221
17.0091 | -105.2217
17.0128 | -105.2189
17.0164 | -105.2165
17.02 | -105.2151
17.0237 | -105.2124
17.0273 | -105.207
17.031 | -105.1994
17.0346 | -105.1928
17.0383 | -105.1889
17.0419 | -105.1828
17.0455 | -105.1767
17.0492 | -105.1702
17.0528 | -105.1623
17.0565 | -105.1521
17.0601 | -105.1415
17.0638 | -105.131
17.0674 | -105.1197
17.071 | -105.1113
17.0747 | -105.105
17.0783 | -105.0968
17.082 | -105.0886
17.0856 | -105.0787
17.0893 | -105.0718
17.0929 | -105.0658
17.0965 | -105.0631
17.1002 | -105.0601
17.1038 | -105.058
17.1075 | -105.0554
17.1111 | -105.055
17.1148 | -105.0556
17.1184 | -105.0579
17.122 | -105.0603
17.1257 | -105.0645
17.1293 | -105.0679
17.133 | -105.0734
17.1366 | -105.08
17.1403 | -105.0848
17.1439 | -105.0902
17.1475 | -105.0967
17.1512 | -105.1062
17.1548 | -105.1145
17.1585 | -105.1208
17.1621 | -105.1298
17.1658 | -105.1402
17.1694 | -105.1522
17.173 | -105.1654
17.1767 | -105.1764
17.1803 | -105.184
17.184 | -105.1942
17.1876 | -105.2044
17.1913 | -105.214
17.1949 | -105.2243
17.1985 | -105.2365
17.2022 | -105.2478
17.2058 | -105.2588
17.2095 | -105.2682
17.2131 | -105.2741
17.2168 | -105.2799
17.2204 | -105.2859
17.224 | -105.2946
17.2277 | -105.303
17.2313 | -105.3088
17.235 | -105.3131
17.2386 | -105.3124
17.2423 | -105.3096
17.2459 | -105.306
17.2495 | -105.3033
17.2532 | -105.3016
17.2568 | -105.2994
17.2605 | -105.2934
17.2641 | -105.2846
17.2678 | -105.2757
17.2714 | -105.2651
17.275 | -105.253
17.2787 | -105.2411
17.2823 | -105.2285
17.286 | -105.2143
17.2896 | -105.1989
17.2933 | -105.1856
17.2969 | -105.1715
17.3005 | -105.1567
17.3042 | -105.1437
17.3078 | -105.1284
17.3115 | -105.1143
17.3151 | -105.0977
17.3188 | -105.0827
17.3224 | -105.071
17.326 | -105.0591
17.3297 | -105.0491
17.3333 | -105.0361
17.337 | -105.0222
17.3406 | -105.0126
17.3443 | -104.9992
17.3479 | -104.9888
17.3515 | -104.9776
17.3552 | -104.9667
17.3588 | -104.9554
17.3625 | -104.9439
17.3661 | -104.934
17.3698 | -104.925
17.3734 | -104.9142
17.377 | -104.9033
17.3807 | -104.8929
17.3843 | -104.8824
17.388 | -104.873
17.3916 | -104.8633
17.3953 | -104.852
17.3989 | -104.8383
17.4026 | -104.8284
17.4062 | -104.8188
17.4098 | -104.8075
17.4135 | -104.7957
17.4171 | -104.7854
17.4208 | -104.7765
17.4244 | -104.7719
17.4281 | -104.7671
17.4317 | -104.7618
17.4353 | -104.76
17.439 | -104.7581
17.4426 | -104.756
17.4463 | -104.7557
17.4499 | -104.7568
17.4536 | -104.7592
17.4572 | -104.7665
17.4608 | -104.7746
17.4645 | -104.7833
17.4681 | -104.7908
17.4718 | -104.7983
17.4754 | -104.8067
17.4791 | -104.8155
17.4827 | -104.8245
17.4863 | -104.8363
17.49 | -104.8518
17.4936 | -104.8663
17.4973 | -104.8805
17.5009 | -104.8952
17.5046 | -104.9096
17.5082 | -104.9256
17.5118 | -104.9427
17.5155 | -104.9592
17.5191 | -104.972
17.5228 | -104.9851
17.5264 | -104.9964
17.5301 | -105.0067
17.5337 | -105.0189
17.5373 | -105.0311
17.541 | -105.0392
17.5446 | -105.0484
17.5483 | -105.0583
17.5519 | -105.0689
17.5556 | -105.0778
17.5592 | -105.0863
17.5628 | -105.0958
17.5665 | -105.1032
17.5701 | -105.1073
17.5738 | -105.1086
17.5774 | -105.1092
17.5811 | -105.1101
17.5847 | -105.1077
17.5883 | -105.1024
17.592 | -105.0945
17.5956 | -105.089
17.5993 | -105.0816
17.6029 | -105.0713
17.6066 | -105.0605
17.6102 | -105.0488
17.6138 | -105.0362
17.6175 | -105.0236
17.6211 | -105.0105
17.6248 | -104.9984
17.6284 | -104.9857
17.6321 | -104.9722
17.6357 | -104.9584
17.6393 | -104.9457
17.643 | -104.9316
17.6466 | -104.92
17.6503 | -104.9099
17.6539 | -104.8971
17.6576 | -104.886
17.6612 | -104.8772
17.6648 | -104.869
17.6685 | -104.8596
17.6721 | -104.8497
17.6758 | -104.8407
17.6794 | -104.8365
17.6831 | -104.8316
17.6867 | -104.8291
17.6903 | -104.8286
17.694 | -104.8279
17.6976 | -104.8288
17.7013 | -104.831
17.7049 | -104.8367
17.7086 | -104.8439
17.7122 | -104.8539
17.7158 | -104.8631
17.7195 | -104.8737
17.7231 | -104.8842
17.7268 | -104.8967
17.7304 | -104.9105
17.7341 | -104.9261
17.7377 | -104.9407
17.7413 | -104.9537
17.745 | -104.9673
17.7486 | -104.9818
17.7523 | -104.9962
17.7559 | -105.009
17.7596 | -105.0179
17.7632 | -105.0253
17.7668 | -105.034
17.7705 | -105.0459
17.7741 | -105.0554
17.7778 | -105.0619
17.7814 | -105.0641
17.7851 | -105.0647
17.7887 | -105.0617
17.7923 | -105.0562
17.796 | -105.0512
17.7996 | -105.0444
17.8033 | -105.0342
17.8069 | -105.0228
17.8106 | -105.0088
17.8142 | -104.9923
17.8179 | -104.9765
17.8215 | -104.9625
17.8251 | -104.9499
17.8288 | -104.9352
17.8324 | -104.9197
17.8361 | -104.9047
17.8397 | -104.8882
17.8434 | -104.8735
17.847 | -104.8583
17.8506 | -104.8426
17.8543 | -104.8307
17.8579 | -104.8209
17.8616 | -104.8107
17.8652 | -104.7989
17.8689 | -104.7988
17.8725 | -104.7907
17.8761 | -104.7847
17.8798 | -104.7783
17.8834 | -104.774
17.8871 | -104.7704
17.8907 | -104.7666
17.8944 | -104.7621
17.898 | -104.7614
17.9016 | -104.762
17.9053 | -104.7631
17.9089 | -104.7659
17.9126 | -104.7668
17.9162 | -104.7673
17.9199 | -104.77
17.9235 | -104.7753
17.9271 | -104.7813
17.9308 | -104.7872
17.9344 | -104.7902
17.9381 | -104.7944
17.9417 | -104.7984
17.9454 | -104.803
17.949 | -104.8095
17.9526 | -104.8153
17.9563 | -104.8192
17.9599 | -104.8235
17.9636 | -104.8263
17.9672 | -104.8316
17.9709 | -104.8389
17.9745 | -104.8471
17.9781 | -104.8537
17.9818 | -104.8592
17.9854 | -104.8638
17.9891 | -104.8652
17.9927 | -104.8679
17.9964 | -104.873
18 | -104.8774
18.0036 | -104.8791
18.0073 | -104.8799
18.0109 | -104.8842
18.0146 | -104.8912
18.0182 | -104.8964
18.0219 | -104.8989
18.0255 | -104.8989
18.0291 | -104.8986
18.0328 | -104.8985
18.0364 | -104.8987
18.0401 | -104.8968
18.0437 | -104.8952
18.0474 | -104.895
18.051 | -104.8911
18.0546 | -104.8869
18.0583 | -104.884
18.0619 | -104.8816
18.0656 | -104.8767
18.0692 | -104.8702
18.0729 | -104.8662
18.0765 | -104.8606
18.0801 | -104.8574
18.0838 | -104.8529
18.0874 | -104.8464
18.0911 | -104.8431
18.0947 | -104.8416
18.0984 | -104.8388
18.102 | -104.8343
18.1056 | -104.8343
18.1093 | -104.833
18.1129 | -104.8308
18.1166 | -104.8304
18.1202 | -104.8326
18.1239 | -104.8317
18.1275 | -104.8325
18.1311 | -104.8332
18.1348 | -104.8357
18.1384 | -104.8389
18.1421 | -104.841
18.1457 | -104.8431
18.1494 | -104.845
18.153 | -104.8484
18.1566 | -104.8505
18.1603 | -104.8524
18.1639 | -104.8569
18.1676 | -104.8617
18.1712 | -104.8675
18.1749 | -104.8729
18.1785 | -104.8772
18.1821 | -104.8811
18.1858 | -104.8842
18.1894 | -104.8874
18.1931 | -104.894
18.1967 | -104.9019
18.2004 | -104.9103
18.204 | -104.9179
18.2077 | -104.9247
18.2113 | -104.9315
18.2149 | -104.9418
18.2186 | -104.951
18.2222 | -104.9595
18.2259 | -104.9652
18.2295 | -104.9723
18.2332 | -104.9798
18.2368 | -104.9892
18.2404 | -104.9966
18.2441 | -105.0046
18.2477 | -105.0141
18.2514 | -105.0236
18.255 | -105.0291
18.2587 | -105.0337
18.2623 | -105.0398
18.2659 | -105.0469
18.2696 | -105.0517
18.2732 | -105.0541
18.2769 | -105.0564
18.2805 | -105.0586
18.2842 | -105.0606
18.2878 | -105.0596
18.2914 | -105.0588
18.2951 | -105.0612
18.2987 | -105.0638
18.3024 | -105.0653
18.306 | -105.0654
18.3097 | -105.0651
18.3133 | -105.0633
18.3169 | -105.062
18.3206 | -105.0614
18.3242 | -105.0621
18.3279 | -105.0635
18.3315 | -105.059
18.3352 | -105.0551
18.3388 | -105.0526
18.3424 | -105.0489
18.3461 | -105.0437
18.3497 | -105.0365
18.3534 | -105.0302
18.357 | -105.022
18.3607 | -105.0133
18.3643 | -105.0027
18.3679 | -104.993
18.3716 | -104.9852
18.3752 | -104.9775
18.3789 | -104.968
18.3825 | -104.9577
18.3862 | -104.9428
18.3898 | -104.9252
18.3934 | -104.8962
18.3971 | -104.8808
18.4007 | -104.8655
18.4044 | -104.8493
18.408 | -104.8337
18.4117 | -104.8202
18.4153 | -104.8052
18.4189 | -104.787
18.4226 | -104.7635
18.4262 | -104.7402
18.4299 | -104.716
18.4335 | -104.6925
18.4372 | -104.6672
18.4408 | -104.6422
18.4444 | -104.6172
18.4481 | -104.5938
18.4517 | -104.5685
18.4554 | -104.5412
18.459 | -104.512
18.4627 | -104.4811
18.4663 | -104.451
18.4699 | -104.4238
18.4736 | -104.394
18.4772 | -104.3638
18.4809 | -104.334
18.4845 | -104.305
18.4882 | -104.2778
18.4918 | -104.2477
18.4954 | -104.2182
18.4991 | -104.1889
18.5027 | -104.1599
18.5064 | -104.1292
18.51 | -104.1001
18.5137 | -104.0706
18.5173 | -104.0423
18.5209 | -104.015
18.5246 | -103.9891
18.5282 | -103.9626
18.5319 | -103.9366
18.5355 | -103.9136
18.5392 | -103.8919
18.5428 | -103.8709
18.5464 | -103.8513
18.5501 | -103.8353
18.5537 | -103.8176
18.5574 | -103.7989
18.561 | -103.7815
18.5647 | -103.7646
18.5683 | -103.7476
18.5719 | -103.7321
18.5756 | -103.7212
18.5792 | -103.7098
18.5829 | -103.6988
18.5865 | -103.6891
18.5902 | -103.6805
18.5938 | -103.6736
18.5974 | -103.6677
18.6011 | -103.6605
18.6047 | -103.657
18.6084 | -103.6576
18.612 | -103.6575
18.6157 | -103.6565
18.6193 | -103.6559
18.623 | -103.6567
18.6266 | -103.6575
18.6302 | -103.6573
18.6339 | -103.6573
18.6375 | -103.6584
18.6412 | -103.6603
18.6448 | -103.663
18.6485 | -103.6683
18.6521 | -103.674
18.6557 | -103.6792
18.6594 | -103.6827
18.663 | -103.6822
18.6667 | -103.6821
18.6703 | -103.6813
18.674 | -103.6807
18.6776 | -103.6789
18.6812 | -103.6756
18.6849 | -103.6732
18.6885 | -103.6721
18.6922 | -103.6684
18.6958 | -103.6663
18.6995 | -103.6629
18.7031 | -103.6592
18.7067 | -103.6556
18.7104 | -103.6533
18.714 | -103.6511
18.7177 | -103.6489
18.7213 | -103.6464
18.725 | -103.6463
18.7286 | -103.6471
18.7322 | -103.6485
18.7359 | -103.6512
18.7395 | -103.6597
18.7432 | -103.6675
18.7468 | -103.6737
18.7505 | -103.6821
18.7541 | -103.6904
18.7577 | -103.7012
18.7614 | -103.7134
18.765 | -103.7268
18.7687 | -103.7394
18.7723 | -103.752
18.776 | -103.7658
18.7796 | -103.7804
18.7832 | -103.7956
18.7869 | -103.8109
18.7905 | -103.8237
18.7942 | -103.8377
18.7978 | -103.854
18.8015 | -103.8685
18.8051 | -103.882
18.8087 | -103.8955
18.8124 | -103.9081
18.816 | -103.9186
18.8197 | -103.9307
18.8233 | -103.942
18.827 | -103.9521
18.8306 | -103.9624
18.8342 | -103.9716
18.8379 | -103.9774
18.8415 | -103.9801
18.8452 | -103.9829
18.8488 | -103.9859
18.8525 | -103.9894
18.8561 | -103.9928
18.8597 | -103.9958
18.8634 | -103.9981
18.867 | -103.9994
18.8707 | -103.9978
18.8743 | -103.9953
18.878 | -103.9946
18.8816 | -103.9949
18.8852 | -103.9966
18.8889 | -103.9957
18.8925 | -103.9939
18.8962 | -103.9935
18.8998 | -103.9916
18.9035 | -103.9902
18.9071 | -103.9877
18.9107 | -103.9841
18.9144 | -103.9785
18.918 | -103.9756
18.9217 | -103.9722
18.9253 | -103.9675
18.929 | -103.9612
18.9326 | -103.9525
18.9362 | -103.9467
18.9399 | -103.9435
18.9435 | -103.9386
18.9472 | -103.9335
18.9508 | -103.928
18.9545 | -103.925
18.9581 | -103.9236
18.9617 | -103.9201
18.9654 | -103.914
18.969 | -103.9094
18.9727 | -103.9054
18.9763 | -103.9007
18.98 | -103.8944
18.9836 | -103.8877
18.9872 | -103.8799
18.9909 | -103.8731
18.9945 | -103.865
18.9982 | -103.8552
19.0018 | -103.8461
19.0055 | -103.8347
19.0091 | -103.821
19.0128 | -103.8093
19.0164 | -103.8017
19.02 | -103.7929
19.0237 | -103.7826
19.0273 | -103.7719
19.031 | -103.7613
19.0346 | -103.7532
19.0383 | -103.7469
19.0419 | -103.7384
19.0455 | -103.7283
19.0492 | -103.7183
19.0528 | -103.707
19.0565 | -103.696
19.0601 | -103.6883
19.0638 | -103.6813
19.0674 | -103.6714
19.071 | -103.6632
19.0747 | -103.6583
19.0783 | -103.6551
19.082 | -103.6517
19.0856 | -103.6486
19.0893 | -103.6458
19.0929 | -103.6434
19.0965 | -103.6424
19.1002 | -103.6394
19.1038 | -103.6362
19.1075 | -103.6361
19.1111 | -103.6367
19.1148 | -103.6368
19.1184 | -103.6372
19.122 | -103.6359
19.1257 | -103.6357
19.1293 | -103.6372
19.133 | -103.6376
19.1366 | -103.6359
19.1403 | -103.6311
19.1439 | -103.6291
19.1475 | -103.6277
19.1512 | -103.6242
19.1548 | -103.6196
19.1585 | -103.6162
19.1621 | -103.6141
19.1658 | -103.612
19.1694 | -103.6107
19.173 | -103.608
19.1767 | -103.6042
19.1803 | -103.6009
19.184 | -103.5971
19.1876 | -103.5941
19.1913 | -103.5887
19.1949 | -103.5836
19.1985 | -103.5808
19.2022 | -103.5803
19.2058 | -103.5799
19.2095 | -103.5764
19.2131 | -103.5711
19.2168 | -103.5661
19.2204 | -103.5619
19.224 | -103.5589
19.2277 | -103.5544
19.2313 | -103.5498
19.235 | -103.5476
19.2386 | -103.546
19.2423 | -103.5447
19.2459 | -103.543
19.2495 | -103.5433
19.2532 | -103.5444
19.2568 | -103.5461
19.2605 | -103.5508
19.2641 | -103.5565
19.2678 | -103.5635
19.2714 | -103.5704
19.275 | -103.5757
19.2787 | -103.5839
19.2823 | -103.5913
19.286 | -103.5999
19.2896 | -103.6106
19.2933 | -103.622
19.2969 | -103.6353
19.3005 | -103.6506
19.3042 | -103.6662
19.3078 | -103.681
19.3115 | -103.6975
19.3151 | -103.7169
19.3188 | -103.7375
19.3224 | -103.7606
19.326 | -103.7853
19.3297 | -103.8097
19.3333 | -103.8339
19.337 | -103.8572
19.3406 | -103.8817
19.3443 | -103.9068
19.3479 | -103.9326
19.3515 | -103.9569
19.3552 | -103.9823
19.3588 | -104.0078
19.3625 | -104.0328
19.3661 | -104.0591
19.3698 | -104.0844
19.3734 | -104.1123
19.377 | -104.1403
19.3807 | -104.1655
19.3843 | -104.1889
19.388 | -104.2107
19.3916 | -104.2305
19.3953 | -104.2493
19.3989 | -104.2671
19.4026 | -104.2845
19.4062 | -104.3022
19.4098 | -104.3217
19.4135 | -104.3402
19.4171 | -104.3533
19.4208 | -104.3648
19.4244 | -104.3778
19.4281 | -104.3911
19.4317 | -104.403
19.4353 | -104.413
19.439 | -104.4202
19.4426 | -104.416
19.4463 | -104.4191
19.4499 | -104.4216
19.4536 | -104.4232
19.4572 | -104.4226
19.4608 | -104.4206
19.4645 | -104.415
19.4681 | -104.4107
19.4718 | -104.4052
19.4754 | -104.3967
19.4791 | -104.3843
19.4827 | -104.3712
19.4863 | -104.3578
19.49 | -104.3447
19.4936 | -104.3309
19.4973 | -104.3182
19.5009 | -104.3047
19.5046 | -104.2888
19.5082 | -104.2739
19.5118 | -104.2583
19.5155 | -104.2436
19.5191 | -104.2278
19.5228 | -104.2142
19.5264 | -104.1983
19.5301 | -104.1827
19.5337 | -104.1668
19.5373 | -104.1526
19.541 | -104.1384
19.5446 | -104.1264
19.5483 | -104.1175
19.5519 | -104.109
19.5556 | -104.1006
19.5592 | -104.0951
19.5628 | -104.0894
19.5665 | -104.0813
19.5701 | -104.0748
19.5738 | -104.0707
19.5774 | -104.0674
19.5811 | -104.0656
19.5847 | -104.0662
19.5883 | -104.0684
19.592 | -104.0723
19.5956 | -104.0746
19.5993 | -104.0767
19.6029 | -104.0805
19.6066 | -104.0846
19.6102 | -104.0911
19.6138 | -104.0977
19.6175 | -104.105
19.6211 | -104.1134
19.6248 | -104.1211
19.6284 | -104.1289
19.6321 | -104.1386
19.6357 | -104.1482
19.6393 | -104.1574
19.643 | -104.1648
19.6466 | -104.1691
19.6503 | -104.1725
19.6539 | -104.1776
19.6576 | -104.1837
19.6612 | -104.1889
19.6648 | -104.1936
19.6685 | -104.1968
19.6721 | -104.1996
19.6758 | -104.1998
19.6794 | -104.1979
19.6831 | -104.1946
19.6867 | -104.1899
19.6903 | -104.1876
19.694 | -104.1812
19.6976 | -104.1733
19.7013 | -104.1642
19.7049 | -104.1538
19.7086 | -104.141
19.7122 | -104.1259
19.7158 | -104.1114
19.7195 | -104.098
19.7231 | -104.0867
19.7268 | -104.0745
19.7304 | -104.061
19.7341 | -104.0482
19.7377 | -104.0361
19.7413 | -104.0221
19.745 | -104.0081
19.7486 | -103.9937
19.7523 | -103.978
19.7559 | -103.9614
19.7596 | -103.9478
19.7632 | -103.9344
19.7668 | -103.9216
19.7705 | -103.9094
19.7741 | -103.8987
19.7778 | -103.8872
19.7814 | -103.8767
19.7851 | -103.868
19.7887 | -103.8606
19.7923 | -103.8534
19.796 | -103.8462
19.7996 | -103.8407
19.8033 | -103.8344
19.8069 | -103.8289
19.8106 | -103.825
19.8142 | -103.823
19.8179 | -103.8201
19.8215 | -103.8176
19.8251 | -103.8139
19.8288 | -103.8104
19.8324 | -103.8079
19.8361 | -103.8054
19.8397 | -103.8026
19.8434 | -103.8006
19.847 | -103.7976
19.8506 | -103.793
19.8543 | -103.7899
19.8579 | -103.787
19.8616 | -103.7827
19.8652 | -103.7777
19.8689 | -103.7715
19.8725 | -103.7661
19.8761 | -103.7619
19.8798 | -103.7563
19.8834 | -103.7503
19.8871 | -103.744
19.8907 | -103.7341
19.8944 | -103.7261
19.898 | -103.7203
19.9016 | -103.7141
19.9053 | -103.7063
19.9089 | -103.6992
19.9126 | -103.6897
19.9162 | -103.6806
19.9199 | -103.672
19.9235 | -103.664
19.9271 | -103.659
19.9308 | -103.6536
19.9344 | -103.6455
19.9381 | -103.6354
19.9417 | -103.6239
19.9454 | -103.6126
19.949 | -103.6038
19.9526 | -103.5958
19.9563 | -103.5856
19.9599 | -103.574
19.9636 | -103.5651
19.9672 | -103.5573
19.9709 | -103.549
19.9745 | -103.5412
19.9781 | -103.5341
19.9818 | -103.5266
19.9854 | -103.52
19.9891 | -103.5113
19.9927 | -103.5023
19.9964 | -103.4958
20 | -103.4878
20.0036 | -103.4778
20.0073 | -103.4634
20.0109 | -103.4509
20.0146 | -103.4423
20.0182 | -103.4361
20.0219 | -103.4302
20.0255 | -103.4245
20.0291 | -103.4161
20.0328 | -103.4079
20.0364 | -103.4007
20.0401 | -103.3942
20.0437 | -103.3867
20.0474 | -103.3775
20.051 | -103.367
20.0546 | -103.3567
20.0583 | -103.346
20.0619 | -103.3334
20.0656 | -103.3201
20.0692 | -103.3094
20.0729 | -103.2974
20.0765 | -103.287
20.0801 | -103.2767
20.0838 | -103.2648
20.0874 | -103.2525
20.0911 | -103.2411
20.0947 | -103.2297
20.0984 | -103.2176
20.102 | -103.2071
20.1056 | -103.1974
20.1093 | -103.1879
20.1129 | -103.1784
20.1166 | -103.1692
20.1202 | -103.162
20.1239 | -103.156
20.1275 | -103.1495
20.1311 | -103.1446
20.1348 | -103.1411
20.1384 | -103.1374
20.1421 | -103.1347
20.1457 | -103.1332
20.1494 | -103.133
20.153 | -103.1319
20.1566 | -103.13
20.1603 | -103.1292
20.1639 | -103.1314
20.1676 | -103.1362
20.1712 | -103.1427
20.1749 | -103.1495
20.1785 | -103.1563
20.1821 | -103.1626
20.1858 | -103.1682
20.1894 | -103.1752
20.1931 | -103.1833
20.1967 | -103.1902
20.2004 | -103.1954
20.204 | -103.2026
20.2077 | -103.208
20.2113 | -103.2125
20.2149 | -103.2146
20.2186 | -103.2169
20.2222 | -103.222
20.2259 | -103.2277
20.2295 | -103.2338
20.2332 | -103.2363
20.2368 | -103.2367
20.2404 | -103.2374
20.2441 | -103.2369
20.2477 | -103.2344
20.2514 | -103.232
20.255 | -103.2306
20.2587 | -103.2271
20.2623 | -103.2215
20.2659 | -103.2183
20.2696 | -103.2133
20.2732 | -103.2043
20.2769 | -103.1981
20.2805 | -103.1915
20.2842 | -103.1843
20.2878 | -103.1775
20.2914 | -103.1707
20.2951 | -103.1611
20.2987 | -103.1504
20.3024 | -103.1378
20.306 | -103.1256
20.3097 | -103.1128
20.3133 | -103.101
20.3169 | -103.0928
20.3206 | -103.0827
20.3242 | -103.0712
20.3279 | -103.0611
20.3315 | -103.0502
20.3352 | -103.0402
20.3388 | -103.0316
20.3424 | -103.0205
20.3461 | -103.0104
20.3497 | -103.0012
20.3534 | -102.9925
20.357 | -102.9844
20.3607 | -102.9773
20.3643 | -102.9705
20.3679 | -102.966
20.3716 | -102.9617
20.3752 | -102.958
20.3789 | -102.9548
20.3825 | -102.9535
20.3862 | -102.9517
20.3898 | -102.9481
20.3934 | -102.947
20.3971 | -102.9462
20.4007 | -102.9476
20.4044 | -102.9492
20.408 | -102.947
20.4117 | -102.9441
20.4153 | -102.9415
20.4189 | -102.9394
20.4226 | -102.9424
20.4262 | -102.9419
20.4299 | -102.94
20.4335 | -102.9402
20.4372 | -102.9401
20.4408 | -102.9416
20.4444 | -102.9426
20.4481 | -102.9435
20.4517 | -102.9434
20.4554 | -102.9406
20.459 | -102.9346
20.4627 | -102.9305
20.4663 | -102.9257
20.4699 | -102.9187
20.4736 | -102.9108
20.4772 | -102.9007
20.4809 | -102.892
20.4845 | -102.8857
20.4882 | -102.8792
20.4918 | -102.8717
20.4954 | -102.8566
20.4991 | -102.8449
20.5027 | -102.8319
20.5064 | -102.8186
20.51 | -102.8077
20.5137 | -102.7942
20.5173 | -102.7822
20.5209 | -102.7668
20.5246 | -102.7513
20.5282 | -102.7385
20.5319 | -102.7274
20.5355 | -102.7176
20.5392 | -102.7061
20.5428 | -102.6932
20.5464 | -102.6776
20.5501 | -102.6604
20.5537 | -102.6451
20.5574 | -102.6265
20.561 | -102.6091
20.5647 | -102.5939
20.5683 | -102.5805
20.5719 | -102.5677
20.5756 | -102.5505
20.5792 | -102.533
20.5829 | -102.5171
20.5865 | -102.502
20.5902 | -102.4852
20.5938 | -102.467
20.5974 | -102.4484
20.6011 | -102.4318
20.6047 | -102.4134
20.6084 | -102.3954
20.612 | -102.3782
20.6157 | -102.3599
20.6193 | -102.3384
20.623 | -102.3168
20.6266 | -102.2975
20.6302 | -102.2774
20.6339 | -102.2556
20.6375 | -102.2327
20.6412 | -102.2128
20.6448 | -102.1922
20.6485 | -102.1706
20.6521 | -102.1488
20.6557 | -102.1261
20.6594 | -102.1025
20.663 | -102.0813
20.6667 | -102.0611
20.6703 | -102.0389
20.674 | -102.0154
20.6776 | -101.9938
20.6812 | -101.9742
20.6849 | -101.9532
20.6885 | -101.9334
20.6922 | -101.9118
20.6958 | -101.893
20.6995 | -101.8788
20.7031 | -101.8654
20.7067 | -101.8515
20.7104 | -101.8374
20.714 | -101.8231
20.7177 | -101.8131
20.7213 | -101.8038
20.725 | -101.7951
20.7286 | -101.7885
20.7322 | -101.7842
20.7359 | -101.7832
20.7395 | -101.7825
20.7432 | -101.7829
20.7468 | -101.7866
20.7505 | -101.7948
20.7541 | -101.8037
20.7577 | -101.8144
20.7614 | -101.8289
20.765 | -101.8454
20.7687 | -101.8644
20.7723 | -101.8837
20.776 | -101.9055
20.7796 | -101.9301
20.7832 | -101.9554
20.7869 | -101.9836
20.7905 | -102.0155
20.7942 | -102.0492
20.7978 | -102.0854
20.8015 | -102.1244
20.8051 | -102.1614
20.8087 | -102.1994
20.8124 | -102.2383
20.816 | -102.281
20.8197 | -102.3243
20.8233 | -102.3671
20.827 | -102.4105
20.8306 | -102.4537
20.8342 | -102.4986
20.8379 | -102.5453
20.8415 | -102.595
20.8452 | -102.6479
20.8488 | -102.6976
20.8525 | -102.7482
20.8561 | -102.7991
20.8597 | -102.8487
20.8634 | -102.8944
20.867 | -102.9392
20.8707 | -102.9856
20.8743 | -103.0324
20.878 | -103.0777
20.8816 | -103.1217
20.8852 | -103.1644
20.8889 | -103.2066
20.8925 | -103.2426
20.8962 | -103.2766
20.8998 | -103.3118
20.9035 | -103.3417
20.9071 | -103.3706
20.9107 | -103.4007
20.9144 | -103.4328
20.918 | -103.4614
20.9217 | -103.4846
20.9253 | -103.5059
20.929 | -103.5296
20.9326 | -103.5509
20.9362 | -103.5644
20.9399 | -103.577
20.9435 | -103.5874
20.9472 | -103.596
20.9508 | -103.6062
20.9545 | -103.6145
20.9581 | -103.6194
20.9617 | -103.624
20.9654 | -103.63
20.969 | -103.6354
20.9727 | -103.6387
20.9763 | -103.6398
20.98 | -103.641
20.9836 | -103.643
20.9872 | -103.6463
20.9909 | -103.6462
20.9945 | -103.6441
20.9982 | -103.6423
21.0018 | -103.6433
21.0055 | -103.6462
21.0091 | -103.6462
21.0128 | -103.6464
21.0164 | -103.6483
21.02 | -103.6557
21.0237 | -103.6592
21.0273 | -103.6613
21.031 | -103.6664
21.0346 | -103.6707
21.0383 | -103.6761
21.0419 | -103.6845
21.0455 | -103.6935
21.0492 | -103.7016
21.0528 | -103.7138
21.0565 | -103.7266
21.0601 | -103.7385
21.0638 | -103.7513
21.0674 | -103.7664
21.071 | -103.7833
21.0747 | -103.7994
21.0783 | -103.8152
21.082 | -103.8304
21.0856 | -103.8467
21.0893 | -103.8655
21.0929 | -103.8829
21.0965 | -103.8999
21.1002 | -103.9163
21.1038 | -103.9332
21.1075 | -103.9503
21.1111 | -103.9659
21.1148 | -103.982
21.1184 | -104.0001
21.122 | -104.017
21.1257 | -104.0321
21.1293 | -104.0474
21.133 | -104.0617
21.1366 | -104.0753
21.1403 | -104.0868
21.1439 | -104.0969
21.1475 | -104.1063
21.1512 | -104.1147
21.1548 | -104.121
21.1585 | -104.1273
21.1621 | -104.1323
21.1658 | -104.1365
21.1694 | -104.1394
21.173 | -104.1423
21.1767 | -104.1443
21.1803 | -104.1484
21.184 | -104.1513
21.1876 | -104.1509
21.1913 | -104.1546
21.1949 | -104.1562
21.1985 | -104.1563
21.2022 | -104.1588
21.2058 | -104.1577
21.2095 | -104.1543
21.2131 | -104.1463
21.2168 | -104.14
21.2204 | -104.1361
21.224 | -104.1308
21.2277 | -104.1236
21.2313 | -104.1181
21.235 | -104.1136
21.2386 | -104.1104
21.2423 | -104.1055
21.2459 | -104.1012
21.2495 | -104.0958
21.2532 | -104.0906
21.2568 | -104.0838
21.2605 | -104.0771
21.2641 | -104.0705
21.2678 | -104.064
21.2714 | -104.0584
21.275 | -104.0544
21.2787 | -104.0486
21.2823 | -104.0412
21.286 | -104.0354
21.2896 | -104.029
21.2933 | -104.0232
21.2969 | -104.0198
21.3005 | -104.0161
21.3042 | -104.0122
21.3078 | -104.0095
21.3115 | -104.0052
21.3151 | -104.0027
21.3188 | -104.0002
21.3224 | -103.9976
21.326 | -103.9949
21.3297 | -103.9936
21.3333 | -103.9917
21.337 | -103.9901
21.3406 | -103.9903
21.3443 | -103.9881
21.3479 | -103.987
21.3515 | -103.9854
21.3552 | -103.9805
21.3588 | -103.9757
21.3625 | -103.974
21.3661 | -103.9733
21.3698 | -103.9684
21.3734 | -103.9625
21.377 | -103.9561
21.3807 | -103.9504
21.3843 | -103.9449
21.388 | -103.9394
21.3916 | -103.9318
21.3953 | -103.9236
21.3989 | -103.9171
21.4026 | -103.9095
21.4062 | -103.9014
21.4098 | -103.8912
21.4135 | -103.8806
21.4171 | -103.8664
21.4208 | -103.8471
21.4244 | -103.8284
21.4281 | -103.8095
21.4317 | -103.7923
21.4353 | -103.7713
21.439 | -103.7453
21.4426 | -103.7188
21.4463 | -103.6923
21.4499 | -103.6676
21.4536 | -103.6412
21.4572 | -103.613
21.4608 | -103.583
21.4645 | -103.5505
21.4681 | -103.5189
21.4718 | -103.4877
21.4754 | -103.452
21.4791 | -103.4137
21.4827 | -103.3758
21.4863 | -103.3355
21.49 | -103.2967
21.4936 | -103.2567
21.4973 | -103.2163
21.5009 | -103.177
21.5046 | -103.1341
21.5082 | -103.0927
21.5118 | -103.0486
21.5155 | -103.0044
21.5191 | -102.9624
21.5228 | -102.9209
21.5264 | -102.8791
21.5301 | -102.837
21.5337 | -102.7949
21.5373 | -102.7521
21.541 | -102.7087
21.5446 | -102.6676
21.5483 | -102.6328
21.5519 | -102.5933
21.5556 | -102.5558
21.5592 | -102.5209
21.5628 | -102.487
21.5665 | -102.456
21.5701 | -102.4263
21.5738 | -102.3961
21.5774 | -102.3656
21.5811 | -102.3339
21.5847 | -102.3046
21.5883 | -102.2785
21.592 | -102.255
21.5956 | -102.2338
21.5993 | -102.2134
21.6029 | -102.1942
21.6066 | -102.1766
21.6102 | -102.1605
21.6138 | -102.1453
21.6175 | -102.1314
21.6211 | -102.1183
21.6248 | -102.1075
21.6284 | -102.0987
21.6321 | -102.0915
21.6357 | -102.0853
21.6393 | -102.0786
21.643 | -102.0744
21.6466 | -102.0711
21.6503 | -102.0702
21.6539 | -102.07
21.6576 | -102.0705
21.6612 | -102.0727
21.6648 | -102.0763
21.6685 | -102.0807
21.6721 | -102.0866
21.6758 | -102.0951
21.6794 | -102.103
21.6831 | -102.1117
21.6867 | -102.1219
21.6903 | -102.133
21.694 | -102.1404
21.6976 | -102.1501
21.7013 | -102.16
21.7049 | -102.1715
21.7086 | -102.1845
21.7122 | -102.1962
21.7158 | -102.21
21.7195 | -102.2257
21.7231 | -102.2421
21.7268 | -102.261
21.7304 | -102.2776
21.7341 | -102.294
21.7377 | -102.3088
21.7413 | -102.3257
21.745 | -102.3425
21.7486 | -102.3589
21.7523 | -102.3751
21.7559 | -102.3897
21.7596 | -102.406
21.7632 | -102.4222
21.7668 | -102.4367
21.7705 | -102.4521
21.7741 | -102.4656
21.7778 | -102.4787
21.7814 | -102.4919
21.7851 | -102.5042
21.7887 | -102.5149
21.7923 | -102.5254
21.796 | -102.5348
21.7996 | -102.5446
21.8033 | -102.5542
21.8069 | -102.5614
21.8106 | -102.5679
21.8142 | -102.573
21.8179 | -102.5811
21.8215 | -102.5892
21.8251 | -102.5929
21.8288 | -102.5935
21.8324 | -102.5965
21.8361 | -102.5993
21.8397 | -102.6002
21.8434 | -102.6036
21.847 | -102.6065
21.8506 | -102.6066
21.8543 | -102.6068
21.8579 | -102.6049
21.8616 | -102.6054
21.8652 | -102.606
21.8689 | -102.6041
21.8725 | -102.603
21.8761 | -102.5996
21.8798 | -102.596
21.8834 | -102.5953
21.8871 | -102.5951
21.8907 | -102.5966
21.8944 | -102.5963
21.898 | -102.5953
21.9016 | -102.5928
21.9053 | -102.5894
21.9089 | -102.586
21.9126 | -102.5804
21.9162 | -102.5749
21.9199 | -102.5703
21.9235 | -102.5666
21.9271 | -102.5598
21.9308 | -102.5528
21.9344 | -102.5484
21.9381 | -102.5432
21.9417 | -102.5375
21.9454 | -102.5319
21.949 | -102.5251
21.9526 | -102.5187
21.9563 | -102.5106
21.9599 | -102.5024
21.9636 | -102.4955
21.9672 | -102.4884
21.9709 | -102.4797
21.9745 | -102.4717
21.9781 | -102.4621
21.9818 | -102.4483
21.9854 | -102.4365
21.9891 | -102.427
21.9927 | -102.418
21.9964 | -102.4083
22 | -102.3988
22.0036 | -102.3918
22.0073 | -102.3827
22.0109 | -102.375
22.0146 | -102.3649
22.0182 | -102.3533
22.0219 | -102.34
22.0255 | -102.3267
22.0291 | -102.3134
22.0328 | -102.3
22.0364 | -102.2892
22.0401 | -102.2799
22.0437 | -102.2702
22.0474 | -102.2594
22.051 | -102.2487
22.0546 | -102.2409
22.0583 | -102.2343
22.0619 | -102.2265
22.0656 | -102.2182
22.0692 | -102.2101
22.0729 | -102.2054
22.0765 | -102.1964
22.0801 | -102.1905
22.0838 | -102.1862
22.0874 | -102.1838
22.0911 | -102.1808
22.0947 | -102.1779
22.0984 | -102.1773
22.102 | -102.1756
22.1056 | -102.1733
22.1093 | -102.1712
22.1129 | -102.1702
22.1166 | -102.1695
22.1202 | -102.1704
22.1239 | -102.1709
22.1275 | -102.172
22.1311 | -102.1726
22.1348 | -102.1708
22.1384 | -102.1694
22.1421 | -102.17
22.1457 | -102.1686
22.1494 | -102.1666
22.153 | -102.1651
22.1566 | -102.1641
22.1603 | -102.1648
22.1639 | -102.1656
22.1676 | -102.1658
22.1712 | -102.1668
22.1749 | -102.1664
22.1785 | -102.1662
22.1821 | -102.1656
22.1858 | -102.164
22.1894 | -102.1623
22.1931 | -102.1617
22.1967 | -102.1621
22.2004 | -102.1628
22.204 | -102.163
22.2077 | -102.1655
22.2113 | -102.1675
22.2149 | -102.169
22.2186 | -102.1721
22.2222 | -102.1763
22.2259 | -102.1835
22.2295 | -102.1917
22.2332 | -102.1983
22.2368 | -102.2038
22.2404 | -102.2145
22.2441 | -102.224
22.2477 | -102.2296
22.2514 | -102.2361
22.255 | -102.2454
22.2587 | -102.2554
22.2623 | -102.2676
22.2659 | -102.2774
22.2696 | -102.2879
22.2732 | -102.2999
22.2769 | -102.3121
22.2805 | -102.3264
22.2842 | -102.3404
22.2878 | -102.3547
22.2914 | -102.3705
22.2951 | -102.3891
22.2987 | -102.4052
22.3024 | -102.4204
22.306 | -102.4358
22.3097 | -102.452
22.3133 | -102.4683
22.3169 | -102.4872
22.3206 | -102.5063
22.3242 | -102.5229
22.3279 | -102.5392
22.3315 | -102.5576
22.3352 | -102.5788
22.3388 | -102.5982
22.3424 | -102.6174
22.3461 | -102.6374
22.3497 | -102.6539
22.3534 | -102.669
22.357 | -102.6848
22.3607 | -102.7018
22.3643 | -102.7181
22.3679 | -102.7334
22.3716 | -102.7457
22.3752 | -102.7564
22.3789 | -102.7658
22.3825 | -102.7739
22.3862 | -102.7833
22.3898 | -102.7935
22.3934 | -102.8004
22.3971 | -102.8046
22.4007 | -102.806
22.4044 | -102.8048
22.408 | -102.8052
22.4117 | -102.8063
22.4153 | -102.8059
22.4189 | -102.8048
22.4226 | -102.8041
22.4262 | -102.8014
22.4299 | -102.7971
22.4335 | -102.7924
22.4372 | -102.7872
22.4408 | -102.7825
22.4444 | -102.7759
22.4481 | -102.7675
22.4517 | -102.7603
22.4554 | -102.7527
22.459 | -102.7422
22.4627 | -102.731
22.4663 | -102.7206
22.4699 | -102.7092
22.4736 | -102.6927
22.4772 | -102.673
22.4809 | -102.6527
22.4845 | -102.6352
22.4882 | -102.6196
22.4918 | -102.603
22.4954 | -102.5846
22.4991 | -102.5682
22.5027 | -102.5535
22.5064 | -102.5386
22.51 | -102.5235
22.5137 | -102.508
22.5173 | -102.4921
22.5209 | -102.4765
22.5246 | -102.4595
22.5282 | -102.4433
22.5319 | -102.4297
22.5355 | -102.4167
22.5392 | -102.4033
22.5428 | -102.3888
22.5464 | -102.3734
22.5501 | -102.3599
22.5537 | -102.3493
22.5574 | -102.3378
22.561 | -102.3259
22.5647 | -102.3154
22.5683 | -102.3043
22.5719 | -102.2915
22.5756 | -102.2806
22.5792 | -102.2703
22.5829 | -102.261
22.5865 | -102.2524
22.5902 | -102.2425
22.5938 | -102.2319
22.5974 | -102.2281
22.6011 | -102.2195
22.6047 | -102.2137
22.6084 | -102.2071
22.612 | -102.2015
22.6157 | -102.1975
22.6193 | -102.1951
22.623 | -102.1935
22.6266 | -102.1932
22.6302 | -102.1931
22.6339 | -102.1941
22.6375 | -102.1963
22.6412 | -102.199
22.6448 | -102.2048
22.6485 | -102.2118
22.6521 | -102.2174
22.6557 | -102.2239
22.6594 | -102.2324
22.663 | -102.242
22.6667 | -102.2519
22.6703 | -102.2635
22.674 | -102.275
22.6776 | -102.2885
22.6812 | -102.3043
22.6849 | -102.3205
22.6885 | -102.3354
22.6922 | -102.3508
22.6958 | -102.3696
22.6995 | -102.39
22.7031 | -102.4096
22.7067 | -102.4275
22.7104 | -102.4452
22.714 | -102.4638
22.7177 | -102.4823
22.7213 | -102.5003
22.725 | -102.5174
22.7286 | -102.5322
22.7322 | -102.5475
22.7359 | -102.5606
22.7395 | -102.5742
22.7432 | -102.5879
22.7468 | -102.6025
22.7505 | -102.616
22.7541 | -102.6267
22.7577 | -102.6348
22.7614 | -102.6427
22.765 | -102.652
22.7687 | -102.6611
22.7723 | -102.6704
22.776 | -102.6777
22.7796 | -102.6817
22.7832 | -102.686
22.7869 | -102.6911
22.7905 | -102.6956
22.7942 | -102.7009
22.7978 | -102.7068
22.8015 | -102.7115
22.8051 | -102.7118
22.8087 | -102.7112
22.8124 | -102.7106
22.816 | -102.7084
22.8197 | -102.7063
22.8233 | -102.7009
22.827 | -102.6943
22.8306 | -102.6868
22.8342 | -102.6759
22.8379 | -102.6659
22.8415 | -102.6565
22.8452 | -102.6453
22.8488 | -102.6347
22.8525 | -102.6211
22.8561 | -102.6062
22.8597 | -102.5926
22.8634 | -102.5767
22.867 | -102.5621
22.8707 | -102.5479
22.8743 | -102.5298
22.878 | -102.5141
22.8816 | -102.5025
22.8852 | -102.4903
22.8889 | -102.4763
22.8925 | -102.4599
22.8962 | -102.4412
22.8998 | -102.4187
22.9035 | -102.3974
22.9071 | -102.3802
22.9107 | -102.3626
22.9144 | -102.3452
22.918 | -102.3251
22.9217 | -102.3038
22.9253 | -102.2825
22.929 | -102.2603
22.9326 | -102.2402
22.9362 | -102.2189
22.9399 | -102.1964
22.9435 | -102.1748
22.9472 | -102.1532
22.9508 | -102.1304
22.9545 | -102.106
22.9581 | -102.0813
22.9617 | -102.058
22.9654 | -102.0366
22.969 | -102.0135
22.9727 | -101.9905
22.9763 | -101.9681
22.98 | -101.9452
22.9836 | -101.9218
22.9872 | -101.8963
22.9909 | -101.8726
22.9945 | -101.8504
22.9982 | -101.8297
23.0018 | -101.8088
23.0055 | -101.7868
23.0091 | -101.7649
23.0128 | -101.745
23.0164 | -101.7264
23.02 | -101.7066
23.0237 | -101.6894
23.0273 | -101.6745
23.031 | -101.6602
23.0346 | -101.6473
23.0383 | -101.6359
23.0419 | -101.625
23.0455 | -101.6156
23.0492 | -101.6072
23.0528 | -101.6011
23.0565 | -101.5924
23.0601 | -101.5843
23.0638 | -101.581
23.0674 | -101.5793
23.071 | -101.5764
23.0747 | -101.5747
23.0783 | -101.5756
23.082 | -101.5783
23.0856 | -101.5827
23.0893 | -101.5871
23.0929 | -101.5909
23.0965 | -101.5963
23.1002 | -101.6033
23.1038 | -101.6124
23.1075 | -101.6218
23.1111 | -101.6304
23.1148 | -101.6406
23.1184 | -101.6517
23.122 | -101.6694
23.1257 | -101.6817
23.1293 | -101.6946
23.133 | -101.7083
23.1366 | -101.7222
23.1403 | -101.7373
23.1439 | -101.7535
23.1475 | -101.769
23.1512 | -101.7877
23.1548 | -101.8057
23.1585 | -101.8232
23.1621 | -101.8437
23.1658 | -101.8646
23.1694 | -101.8843
23.173 | -101.9055
23.1767 | -101.925
23.1803 | -101.9457
23.184 | -101.9676
23.1876 | -101.9883
23.1913 | -102.0092
23.1949 | -102.0307
23.1985 | -102.0504
23.2022 | -102.0708
23.2058 | -102.0888
23.2095 | -102.1055
23.2131 | -102.1228
23.2168 | -102.1408
23.2204 | -102.1617
23.224 | -102.1834
23.2277 | -102.2076
23.2313 | -102.2286
23.235 | -102.2436
23.2386 | -102.2577
23.2423 | -102.2723
23.2459 | -102.2873
23.2495 | -102.3018
23.2532 | -102.3167
23.2568 | -102.3317
23.2605 | -102.3441
23.2641 | -102.3565
23.2678 | -102.3683
23.2714 | -102.3787
23.275 | -102.3876
23.2787 | -102.3962
23.2823 | -102.4076
23.286 | -102.4195
23.2896 | -102.4281
23.2933 | -102.4342
23.2969 | -102.4415
23.3005 | -102.4488
23.3042 | -102.4575
23.3078 | -102.4645
23.3115 | -102.4702
23.3151 | -102.4773
23.3188 | -102.4837
23.3224 | -102.4868
23.326 | -102.4887
23.3297 | -102.4912
23.3333 | -102.4948
23.337 | -102.5009
23.3406 | -102.506
23.3443 | -102.5103
23.3479 | -102.5151
23.3515 | -102.5168
23.3552 | -102.5186
23.3588 | -102.5212
23.3625 | -102.5237
23.3661 | -102.5265
23.3698 | -102.5297
23.3734 | -102.5341
23.377 | -102.5388
23.3807 | -102.5415
23.3843 | -102.5485
23.388 | -102.5579
23.3916 | -102.5632
23.3953 | -102.5673
23.3989 | -102.5726
23.4026 | -102.5775
23.4062 | -102.5813
23.4098 | -102.5856
23.4135 | -102.5912
23.4171 | -102.5945
23.4208 | -102.6004
23.4244 | -102.6066
23.4281 | -102.6122
23.4317 | -102.6179
23.4353 | -102.6237
23.439 | -102.6289
23.4426 | -102.6334
23.4463 | -102.6389
23.4499 | -102.6448
23.4536 | -102.6503
23.4572 | -102.6551
23.4608 | -102.6605
23.4645 | -102.6656
23.4681 | -102.6685
23.4718 | -102.6729
23.4754 | -102.6782
23.4791 | -102.6842
23.4827 | -102.6862
23.4863 | -102.6889
23.49 | -102.6923
23.4936 | -102.6952
23.4973 | -102.7004
23.5009 | -102.7053
23.5046 | -102.7107
23.5082 | -102.7178
23.5118 | -102.7244
23.5155 | -102.7294
23.5191 | -102.7348
23.5228 | -102.7415
23.5264 | -102.7454
23.5301 | -102.7495
23.5337 | -102.7538
23.5373 | -102.7598
23.541 | -102.7649
23.5446 | -102.769
23.5483 | -102.773
23.5519 | -102.7752
23.5556 | -102.7776
23.5592 | -102.7811
23.5628 | -102.7862
23.5665 | -102.7927
23.5701 | -102.8008
23.5738 | -102.8083
23.5774 | -102.8172
23.5811 | -102.826
23.5847 | -102.834
23.5883 | -102.8413
23.592 | -102.8494
23.5956 | -102.8572
23.5993 | -102.8661
23.6029 | -102.8742
23.6066 | -102.8855
23.6102 | -102.8979
23.6138 | -102.9089
23.6175 | -102.9187
23.6211 | -102.9295
23.6248 | -102.94
23.6284 | -102.9521
23.6321 | -102.9654
23.6357 | -102.9775
23.6393 | -102.9926
23.643 | -103.0067
23.6466 | -103.0196
23.6503 | -103.033
23.6539 | -103.0428
23.6576 | -103.0544
23.6612 | -103.0676
23.6648 | -103.0789
23.6685 | -103.0886
23.6721 | -103.0973
23.6758 | -103.1078
23.6794 | -103.1191
23.6831 | -103.1281
23.6867 | -103.1367
23.6903 | -103.1455
23.694 | -103.1537
23.6976 | -103.1626
23.7013 | -103.1692
23.7049 | -103.1749
23.7086 | -103.1806
23.7122 | -103.1849
23.7158 | -103.1892
23.7195 | -103.1924
23.7231 | -103.1948
23.7268 | -103.1956
23.7304 | -103.1963
23.7341 | -103.1978
23.7377 | -103.1986
23.7413 | -103.1991
23.745 | -103.199
23.7486 | -103.1966
23.7523 | -103.1951
23.7559 | -103.1923
23.7596 | -103.1871
23.7632 | -103.1824
23.7668 | -103.1788
23.7705 | -103.1737
23.7741 | -103.1683
23.7778 | -103.1622
23.7814 | -103.1558
23.7851 | -103.1481
23.7887 | -103.1399
23.7923 | -103.1333
23.796 | -103.1277
23.7996 | -103.1214
23.8033 | -103.1176
23.8069 | -103.1134
23.8106 | -103.1089
23.8142 | -103.1044
23.8179 | -103.1035
23.8215 | -103.1042
23.8251 | -103.1048
23.8288 | -103.105
23.8324 | -103.103
23.8361 | -103.1037
23.8397 | -103.1064
23.8434 | -103.1085
23.847 | -103.112
23.8506 | -103.1174
23.8543 | -103.1242
23.8579 | -103.1336
23.8616 | -103.1456
23.8652 | -103.1587
23.8689 | -103.1722
23.8725 | -103.1846
23.8761 | -103.1992
23.8798 | -103.2153
23.8834 | -103.232
23.8871 | -103.25
23.8907 | -103.2682
23.8944 | -103.2851
23.898 | -103.302
23.9016 | -103.3208
23.9053 | -103.3382
23.9089 | -103.3537
23.9126 | -103.3701
23.9162 | -103.3849
23.9199 | -103.4006
23.9235 | -103.4181
23.9271 | -103.435
23.9308 | -103.452
23.9344 | -103.4696
23.9381 | -103.487
23.9417 | -103.4984
23.9454 | -103.5093
23.949 | -103.5171
23.9526 | -103.5277
23.9563 | -103.539
23.9599 | -103.5519
23.9636 | -103.562
23.9672 | -103.5708
23.9709 | -103.5794
23.9745 | -103.588
23.9781 | -103.5974
23.9818 | -103.6049
23.9854 | -103.6117
23.9891 | -103.619
23.9927 | -103.6274
23.9964 | -103.6347
24 | -103.641
24.0036 | -103.6458
24.0073 | -103.6501
24.0109 | -103.6539
24.0146 | -103.6527
24.0182 | -103.6506
24.0219 | -103.6493
24.0255 | -103.6481
24.0291 | -103.6441
24.0328 | -103.6382
24.0364 | -103.6317
24.0401 | -103.6259
24.0437 | -103.619
24.0474 | -103.612
24.051 | -103.6071
24.0546 | -103.6009
24.0583 | -103.5934
24.0619 | -103.5828
24.0656 | -103.5732
24.0692 | -103.563
24.0729 | -103.551
24.0765 | -103.5391
24.0801 | -103.5294
24.0838 | -103.5171
24.0874 | -103.502
24.0911 | -103.4885
24.0947 | -103.4764
24.0984 | -103.4617
24.102 | -103.4484
24.1056 | -103.4362
24.1093 | -103.4255
24.1129 | -103.4149
24.1166 | -103.4038
24.1202 | -103.3915
24.1239 | -103.3791
24.1275 | -103.3695
24.1311 | -103.357
24.1348 | -103.3409
24.1384 | -103.327
24.1421 | -103.3126
24.1457 | -103.2979
24.1494 | -103.2853
24.153 | -103.2724
24.1566 | -103.2599
24.1603 | -103.2461
24.1639 | -103.232
24.1676 | -103.2178
24.1712 | -103.2026
24.1749 | -103.1794
24.1785 | -103.165
24.1821 | -103.1485
24.1858 | -103.1326
24.1894 | -103.1173
24.1931 | -103.1012
24.1967 | -103.0833
24.2004 | -103.0644
24.204 | -103.0429
24.2077 | -103.0222
24.2113 | -103.0038
24.2149 | -102.9836
24.2186 | -102.9585
24.2222 | -102.9215
24.2259 | -102.8532
24.2295 | -102.7685
24.2332 | -102.6729
24.2368 | -102.5769
24.2404 | -102.4883
24.2441 | -102.4133
24.2477 | -102.348
24.2514 | -102.2828
24.255 | -102.2143
24.2587 | -102.1424
24.2623 | -102.0677
24.2659 | -101.9898
24.2696 | -101.9103
24.2732 | -101.829
24.2769 | -101.7434
24.2805 | -101.6549
24.2842 | -101.5658
24.2878 | -101.4736
24.2914 | -101.3781
24.2951 | -101.2786
24.2987 | -101.1788
24.3024 | -101.0779
24.306 | -100.9776
24.3097 | -100.8773
24.3133 | -100.7774
24.3169 | -100.6773
24.3206 | -100.5787
24.3242 | -100.4816
24.3279 | -100.3854
24.3315 | -100.2912
24.3352 | -100.198
24.3388 | -100.1075
24.3424 | -100.0195
24.3461 | -99.9349
24.3497 | -99.8539
24.3534 | -99.7764
24.357 | -99.7044
24.3607 | -99.637
24.3643 | -99.5743
24.3679 | -99.5154
24.3716 | -99.4618
24.3752 | -99.4138
24.3789 | -99.3704
24.3825 | -99.3314
24.3862 | -99.298
24.3898 | -99.2693
24.3934 | -99.246
24.3971 | -99.2282
24.4007 | -99.2164
24.4044 | -99.2112
24.408 | -99.2116
24.4117 | -99.2182
24.4153 | -99.2307
24.4189 | -99.2487
24.4226 | -99.2723
24.4262 | -99.3019
24.4299 | -99.3379
24.4335 | -99.3784
24.4372 | -99.4231
24.4408 | -99.4731
24.4444 | -99.5271
24.4481 | -99.5861
24.4517 | -99.6493
24.4554 | -99.7165
24.459 | -99.7883
24.4627 | -99.864
24.4663 | -99.9417
24.4699 | -100.022
24.4736 | -100.1038
24.4772 | -100.1893
24.4809 | -100.2762
24.4845 | -100.366
24.4882 | -100.4585
24.4918 | -100.5513
24.4954 | -100.6446
24.4991 | -100.7368
24.5027 | -100.829
24.5064 | -100.9209
24.51 | -101.0128
24.5137 | -101.1038
24.5173 | -101.1942
24.5209 | -101.2813
24.5246 | -101.3653
24.5282 | -101.4469
24.5319 | -101.5256
24.5355 | -101.6
24.5392 | -101.671
24.5428 | -101.7401
24.5464 | -101.8051
24.5501 | -101.868
24.5537 | -101.9256
24.5574 | -101.9768
24.561 | -102.0255
24.5647 | -102.0731
24.5683 | -102.1257
24.5719 | -102.1964
24.5756 | -102.2681
24.5792 | -102.3308
24.5829 | -102.3709
24.5865 | -102.3705
24.5902 | -102.3358
24.5938 | -102.2804
24.5974 | -102.2213
24.6011 | -102.1714
24.6047 | -102.1331
24.6084 | -102.095
24.612 | -102.0515
24.6157 | -102.0039
24.6193 | -101.9544
24.623 | -101.9042
24.6266 | -101.8522
24.6302 | -101.7967
24.6339 | -101.7378
24.6375 | -101.6771
24.6412 | -101.6164
24.6448 | -101.553
24.6485 | -101.4875
24.6521 | -101.4184
24.6557 | -101.3471
24.6594 | -101.2762
24.663 | -101.2033
24.6667 | -101.1288
24.6703 | -101.0538
24.674 | -100.9788
24.6776 | -100.903
24.6812 | -100.8289
24.6849 | -100.7543
24.6885 | -100.6805
24.6922 | -100.6071
24.6958 | -100.5343
24.6995 | -100.4682
24.7031 | -100.4026
24.7067 | -100.3377
24.7104 | -100.2738
24.714 | -100.2132
24.7177 | -100.1559
24.7213 | -100.1012
24.725 | -100.0509
24.7286 | -100.0038
24.7322 | -99.9606
24.7359 | -99.9207
24.7395 | -99.8854
24.7432 | -99.8543
24.7468 | -99.8282
24.7505 | -99.8062
24.7541 | -99.7883
24.7577 | -99.7741
24.7614 | -99.7639
24.765 | -99.7595
24.7687 | -99.7598
24.7723 | -99.7644
24.776 | -99.7742
24.7796 | -99.788
24.7832 | -99.8067
24.7869 | -99.8301
24.7905 | -99.8597
24.7942 | -99.8942
24.7978 | -99.9325
24.8015 | -99.9751
24.8051 | -100.0223
24.8087 | -100.0732
24.8124 | -100.1278
24.816 | -100.1872
24.8197 | -100.2505
24.8233 | -100.3189
24.827 | -100.3896
24.8306 | -100.4618
24.8342 | -100.5349
24.8379 | -100.6092
24.8415 | -100.6865
24.8452 | -100.7661
24.8488 | -100.8485
24.8525 | -100.9321
24.8561 | -101.0159
24.8597 | -101.1004
24.8634 | -101.1855
24.867 | -101.2698
24.8707 | -101.3544
24.8743 | -101.4374
24.878 | -101.5194
24.8816 | -101.6001
24.8852 | -101.6806
24.8889 | -101.7605
24.8925 | -101.8379
24.8962 | -101.9132
24.8998 | -101.9848
24.9035 | -102.0545
24.9071 | -102.1231
24.9107 | -102.1874
24.9144 | -102.2473
24.918 | -102.3044
24.9217 | -102.3578
24.9253 | -102.4095
24.929 | -102.4628
24.9326 | -102.5208
24.9362 | -102.5909
24.9399 | -102.6624
24.9435 | -102.7286
24.9472 | -102.7834
24.9508 | -102.8169
24.9545 | -102.8328
24.9581 | -102.8415
24.9617 | -102.8488
24.9654 | -102.8549
24.969 | -102.8589
24.9727 | -102.8612
24.9763 | -102.8602
24.98 | -102.8598
24.9836 | -102.8609
24.9872 | -102.8615
24.9909 | -102.8567
24.9945 | -102.8493
24.9982 | -102.8408
25.0018 | -102.8327
25.0055 | -102.8241
25.0091 | -102.8131
25.0128 | -102.8034
25.0164 | -102.7925
25.02 | -102.7812
25.0237 | -102.7707
25.0273 | -102.7603
25.031 | -102.7479
25.0346 | -102.7359
25.0383 | -102.7249
25.0419 | -102.7146
25.0455 | -102.7039
25.0492 | -102.692
25.0528 | -102.679
25.0565 | -102.6647
25.0601 | -102.6524
25.0638 | -102.6436
25.0674 | -102.634
25.071 | -102.6246
25.0747 | -102.6154
25.0783 | -102.6057
25.082 | -102.5976
25.0856 | -102.5921
25.0893 | -102.5884
25.0929 | -102.5826
25.0965 | -102.5778
25.1002 | -102.5744
25.1038 | -102.5711
25.1075 | -102.5692
25.1111 | -102.5712
25.1148 | -102.5736
25.1184 | -102.5744
25.122 | -102.5765
25.1257 | -102.5805
25.1293 | -102.585
25.133 | -102.5893
25.1366 | -102.5955
25.1403 | -102.6042
25.1439 | -102.6114
25.1475 | -102.6171
25.1512 | -102.6238
25.1548 | -102.6311
25.1585 | -102.6388
25.1621 | -102.6476
25.1658 | -102.6581
25.1694 | -102.671
25.173 | -102.6817
25.1767 | -102.6915
25.1803 | -102.7023
25.184 | -102.7143
25.1876 | -102.7279
25.1913 | -102.74
25.1949 | -102.7529
25.1985 | -102.7661
25.2022 | -102.7787
25.2058 | -102.791
25.2095 | -102.8041
25.2131 | -102.8171
25.2168 | -102.8286
25.2204 | -102.8436
25.224 | -102.857
25.2277 | -102.8685
25.2313 | -102.8816
25.235 | -102.8944
25.2386 | -102.9072
25.2423 | -102.92
25.2459 | -102.9317
25.2495 | -102.9412
25.2532 | -102.952
25.2568 | -102.9649
25.2605 | -102.9775
25.2641 | -102.9876
25.2678 | -102.9996
25.2714 | -103.0112
25.275 | -103.0194
25.2787 | -103.0275
25.2823 | -103.0359
25.286 | -103.0436
25.2896 | -103.0509
25.2933 | -103.0585
25.2969 | -103.0646
25.3005 | -103.07
25.3042 | -103.076
25.3078 | -103.0808
25.3115 | -103.0848
25.3151 | -103.0881
25.3188 | -103.0908
25.3224 | -103.0941
25.326 | -103.0948
25.3297 | -103.0917
25.3333 | -103.0916
25.337 | -103.0899
25.3406 | -103.0873
25.3443 | -103.0851
25.3479 | -103.0838
25.3515 | -103.0808
25.3552 | -103.0768
25.3588 | -103.0726
25.3625 | -103.069
25.3661 | -103.0659
25.3698 | -103.0619
25.3734 | -103.0596
25.377 | -103.0583
25.3807 | -103.0555
25.3843 | -103.052
25.388 | -103.049
25.3916 | -103.0468
25.3953 | -103.0424
25.3989 | -103.0388
25.4026 | -103.0377
25.4062 | -103.0384
25.4098 | -103.0383
25.4135 | -103.0395
25.4171 | -103.0402
25.4208 | -103.039
25.4244 | -103.0375
25.4281 | -103.0376
25.4317 | -103.0381
25.4353 | -103.0385
25.439 | -103.04
25.4426 | -103.0397
25.4463 | -103.0405
25.4499 | -103.044
25.4536 | -103.0465
25.4572 | -103.0499
25.4608 | -103.0552
25.4645 | -103.0608
25.4681 | -103.0656
25.4718 | -103.0705
25.4754 | -103.0767
25.4791 | -103.0823
25.4827 | -103.087
25.4863 | -103.0942
25.49 | -103.1031
25.4936 | -103.1105
25.4973 | -103.1183
25.5009 | -103.1276
25.5046 | -103.1377
25.5082 | -103.1466
25.5118 | -103.1571
25.5155 | -103.1687
25.5191 | -103.1801
25.5228 | -103.1903
25.5264 | -103.1997
25.5301 | -103.212
25.5337 | -103.2243
25.5373 | -103.2356
25.541 | -103.2472
25.5446 | -103.2583
25.5483 | -103.2693
25.5519 | -103.2787
25.5556 | -103.2888
25.5592 | -103.2985
25.5628 | -103.309
25.5665 | -103.3198
25.5701 | -103.3291
25.5738 | -103.3406
25.5774 | -103.3491
25.5811 | -103.3565
25.5847 | -103.3636
25.5883 | -103.3719
25.592 | -103.3787
25.5956 | -103.3869
25.5993 | -103.3939
25.6029 | -103.3986
25.6066 | -103.4038
25.6102 | -103.4107
25.6138 | -103.4156
25.6175 | -103.4183
25.6211 | -103.4188
25.6248 | -103.4198
25.6284 | -103.4172
25.6321 | -103.4147
25.6357 | -103.4158
25.6393 | -103.4172
25.643 | -103.4149
25.6466 | -103.4117
25.6503 | -103.4095
25.6539 | -103.4071
25.6576 | -103.404
25.6612 | -103.3992
25.6648 | -103.3934
25.6685 | -103.3879
25.6721 | -103.3811
25.6758 | -103.3742
25.6794 | -103.3671
25.6831 | -103.3582
25.6867 | -103.3499
25.6903 | -103.346
25.694 | -103.3395
25.6976 | -103.332
25.7013 | -103.3226
25.7049 | -103.3122
25.7086 | -103.3013
25.7122 | -103.2901
25.7158 | -103.2807
25.7195 | -103.2718
25.7231 | -103.2631
25.7268 | -103.2527
25.7304 | -103.2421
25.7341 | -103.2306
25.7377 | -103.2185
25.7413 | -103.2088
25.745 | -103.2015
25.7486 | -103.194
25.7523 | -103.1871
25.7559 | -103.1775
25.7596 | -103.1685
25.7632 | -103.1586
25.7668 | -103.1494
25.7705 | -103.1399
25.7741 | -103.1302
25.7778 | -103.1215
25.7814 | -103.1132
25.7851 | -103.1058
25.7887 | -103.0984
25.7923 | -103.0913
25.796 | -103.0848
25.7996 | -103.0775
25.8033 | -103.0698
25.8069 | -103.0638
25.8106 | -103.058
25.8142 | -103.052
25.8179 | -103.0464
25.8215 | -103.0404
25.8251 | -103.0332
25.8288 | -103.0257
25.8324 | -103.0193
25.8361 | -103.0139
25.8397 | -103.01
25.8434 | -103.0073
25.847 | -103.0038
25.8506 | -102.998
25.8543 | -102.9936
25.8579 | -102.9912
25.8616 | -102.9864
25.8652 | -102.9819
25.8689 | -102.9792
25.8725 | -102.9774
25.8761 | -102.9756
25.8798 | -102.9731
25.8834 | -102.9704
25.8871 | -102.9695
25.8907 | -102.9689
25.8944 | -102.9671
25.898 | -102.9685
25.9016 | -102.9707
25.9053 | -102.9728
25.9089 | -102.9745
25.9126 | -102.9774
25.9162 | -102.9807
25.9199 | -102.9849
25.9235 | -102.991
25.9271 | -102.9983
25.9308 | -103.0054
25.9344 | -103.0143
25.9381 | -103.0232
25.9417 | -103.0326
25.9454 | -103.0428
25.949 | -103.0537
25.9526 | -103.0656
25.9563 | -103.0766
25.9599 | -103.088
25.9636 | -103.1015
25.9672 | -103.1153
25.9709 | -103.1289
25.9745 | -103.144
25.9781 | -103.1591
25.9818 | -103.175
25.9854 | -103.1922
25.9891 | -103.21
25.9927 | -103.2288
25.9964 | -103.2479
26 | -103.2652
26.0036 | -103.2821
26.0073 | -103.2999
26.0109 | -103.3187
26.0146 | -103.3362
26.0182 | -103.3527
26.0219 | -103.3687
26.0255 | -103.3863
26.0291 | -103.4045
26.0328 | -103.4219
26.0364 | -103.4395
26.0401 | -103.4572
26.0437 | -103.4751
26.0474 | -103.4926
26.051 | -103.5104
26.0546 | -103.5262
26.0583 | -103.5411
26.0619 | -103.5564
26.0656 | -103.5712
26.0692 | -103.586
26.0729 | -103.6011
26.0765 | -103.6163
26.0801 | -103.63
26.0838 | -103.6424
26.0874 | -103.653
26.0911 | -103.6636
26.0947 | -103.673
26.0984 | -103.6833
26.102 | -103.6945
26.1056 | -103.7038
26.1093 | -103.7093
26.1129 | -103.7144
26.1166 | -103.7189
26.1202 | -103.7243
26.1239 | -103.7317
26.1275 | -103.7396
26.1311 | -103.7461
26.1348 | -103.7497
26.1384 | -103.754
26.1421 | -103.7581
26.1457 | -103.7601
26.1494 | -103.7626
26.153 | -103.7657
26.1566 | -103.7685
26.1603 | -103.7709
26.1639 | -103.7725
26.1676 | -103.7731
26.1712 | -103.7751
26.1749 | -103.7771
26.1785 | -103.7789
26.1821 | -103.7809
26.1858 | -103.7812
26.1894 | -103.7812
26.1931 | -103.7834
26.1967 | -103.7858
26.2004 | -103.7874
26.204 | -103.7896
26.2077 | -103.791
26.2113 | -103.792
26.2149 | -103.7949
26.2186 | -103.7987
26.2222 | -103.8005
26.2259 | -103.8029
26.2295 | -103.8058
26.2332 | -103.8095
26.2368 | -103.8133
26.2404 | -103.816
26.2441 | -103.8184
26.2477 | -103.8224
26.2514 | -103.8269
26.255 | -103.8309
26.2587 | -103.8355
26.2623 | -103.8411
26.2659 | -103.8445
26.2696 | -103.8485
26.2732 | -103.8529
26.2769 | -103.8581
26.2805 | -103.8637
26.2842 | -103.8704
26.2878 | -103.8784
26.2914 | -103.8857
26.2951 | -103.8923
26.2987 | -103.8977
26.3024 | -103.902
26.306 | -103.9062
26.3097 | -103.9114
26.3133 | -103.9176
26.3169 | -103.9248
26.3206 | -103.9339
26.3242 | -103.9413
26.3279 | -103.9484
26.3315 | -103.9556
26.3352 | -103.962
26.3388 | -103.9685
26.3424 | -103.9755
26.3461 | -103.9826
26.3497 | -103.9892
26.3534 | -103.9947
26.357 | -103.9994
26.3607 | -104.0043
26.3643 | -104.0107
26.3679 | -104.0161
26.3716 | -104.0206
26.3752 | -104.0247
26.3789 | -104.0284
26.3825 | -104.0288
26.3862 | -104.0296
26.3898 | -104.0332
26.3934 | -104.0377
26.3971 | -104.0405
26.4007 | -104.0424
26.4044 | -104.0431
26.408 | -104.0417
26.4117 | -104.0414
26.4153 | -104.0406
26.4189 | -104.0386
26.4226 | -104.0357
26.4262 | -104.0329
26.4299 | -104.0297
26.4335 | -104.0259
26.4372 | -104.022
26.4408 | -104.0169
26.4444 | -104.011
26.4481 | -104.0057
26.4517 | -104
26.4554 | -103.992
26.459 | -103.9855
26.4627 | -103.9783
26.4663 | -103.9709
26.4699 | -103.9638
26.4736 | -103.9546
26.4772 | -103.9441
26.4809 | -103.936
26.4845 | -103.9271
26.4882 | -103.9169
26.4918 | -103.9048
26.4954 | -103.8919
26.4991 | -103.8798
26.5027 | -103.8678
26.5064 | -103.8559
26.51 | -103.8439
26.5137 | -103.8324
26.5173 | -103.822
26.5209 | -103.8098
26.5246 | -103.7972
26.5282 | -103.7843
26.5319 | -103.7715
26.5355 | -103.7591
26.5392 | -103.7474
26.5428 | -103.7355
26.5464 | -103.7237
26.5501 | -103.7119
26.5537 | -103.7011
26.5574 | -103.6919
26.561 | -103.6825
26.5647 | -103.6725
26.5683 | -103.663
26.5719 | -103.6532
26.5756 | -103.6439
26.5792 | -103.6344
26.5829 | -103.6243
26.5865 | -103.6141
26.5902 | -103.6049
26.5938 | -103.5966
26.5974 | -103.5881
26.6011 | -103.5802
26.6047 | -103.5719
26.6084 | -103.5634
26.612 | -103.556
26.6157 | -103.5483
26.6193 | -103.5408
26.623 | -103.5349
26.6266 | -103.529
26.6302 | -103.5231
26.6339 | -103.5174
26.6375 | -103.5122
26.6412 | -103.5064
26.6448 | -103.5008
26.6485 | -103.4949
26.6521 | -103.4881
26.6557 | -103.4813
26.6594 | -103.4755
26.663 | -103.4678
26.6667 | -103.4597
26.6703 | -103.4538
26.674 | -103.4487
26.6776 | -103.4426
26.6812 | -103.4352
26.6849 | -103.4287
26.6885 | -103.4217
26.6922 | -103.4146
26.6958 | -103.4068
26.6995 | -103.3992
26.7031 | -103.3921
26.7067 | -103.3854
26.7104 | -103.3785
26.714 | -103.3715
26.7177 | -103.3636
26.7213 | -103.3548
26.725 | -103.348
26.7286 | -103.3421
26.7322 | -103.3351
26.7359 | -103.329
26.7395 | -103.3219
26.7432 | -103.314
26.7468 | -103.3043
26.7505 | -103.2966
26.7541 | -103.291
26.7577 | -103.2853
26.7614 | -103.2787
26.765 | -103.273
26.7687 | -103.2659
26.7723 | -103.2583
26.776 | -103.2508
26.7796 | -103.2455
26.7832 | -103.2414
26.7869 | -103.236
26.7905 | -103.2314
26.7942 | -103.2289
26.7978 | -103.2254
26.8015 | -103.2238
26.8051 | -103.2195
26.8087 | -103.2164
26.8124 | -103.2135
26.816 | -103.2124
26.8197 | -103.2124
26.8233 | -103.2109
26.827 | -103.2086
26.8306 | -103.207
26.8342 | -103.2072
26.8379 | -103.2064
26.8415 | -103.2068
26.8452 | -103.2096
26.8488 | -103.2106
26.8525 | -103.2104
26.8561 | -103.2115
26.8597 | -103.2121
26.8634 | -103.2138
26.867 | -103.2172
26.8707 | -103.2212
26.8743 | -103.2239
26.878 | -103.2248
26.8816 | -103.2274
26.8852 | -103.2314
26.8889 | -103.2355
26.8925 | -103.2386
26.8962 | -103.2414
26.8998 | -103.2444
26.9035 | -103.2473
26.9071 | -103.2519
26.9107 | -103.2563
26.9144 | -103.2602
26.918 | -103.2644
26.9217 | -103.2686
26.9253 | -103.2725
26.929 | -103.2763
26.9326 | -103.2797
26.9362 | -103.2835
26.9399 | -103.2878
26.9435 | -103.2935
26.9472 | -103.2992
26.9508 | -103.3048
26.9545 | -103.3104
26.9581 | -103.3167
26.9617 | -103.323
26.9654 | -103.3299
26.969 | -103.3372
26.9727 | -103.3447
26.9763 | -103.3516
26.98 | -103.3589
26.9836 | -103.3663
26.9872 | -103.3742
26.9909 | -103.383
26.9945 | -103.3911
26.9982 | -103.398
27.0018 | -103.4058
27.0055 | -103.4137
27.0091 | -103.4219
27.0128 | -103.4309
27.0164 | -103.439
27.02 | -103.4466
27.0237 | -103.4543
27.0273 | -103.4634
27.031 | -103.4734
27.0346 | -103.4825
27.0383 | -103.4914
27.0419 | -103.501
27.0455 | -103.511
27.0492 | -103.5189
27.0528 | -103.5262
27.0565 | -103.5347
27.0601 | -103.5444
27.0638 | -103.5524
27.0674 | -103.5607
27.071 | -103.568
27.0747 | -103.5752
27.0783 | -103.5841
27.082 | -103.5917
27.0856 | -103.5971
27.0893 | -103.6024
27.0929 | -103.6062
27.0965 | -103.6098
27.1002 | -103.6147
27.1038 | -103.6186
27.1075 | -103.6214
27.1111 | -103.6255
27.1148 | -103.6276
27.1184 | -103.6282
27.122 | -103.6308
27.1257 | -103.6355
27.1293 | -103.6359
27.133 | -103.6348
27.1366 | -103.6354
27.1403 | -103.635
27.1439 | -103.6309
27.1475 | -103.6269
27.1512 | -103.624
27.1548 | -103.6215
27.1585 | -103.6184
27.1621 | -103.6143
27.1658 | -103.6109
27.1694 | -103.6076
27.173 | -103.6026
27.1767 | -103.5985
27.1803 | -103.5955
27.184 | -103.5927
27.1876 | -103.589
27.1913 | -103.5848
27.1949 | -103.5799
27.1985 | -103.5744
27.2022 | -103.5684
27.2058 | -103.5613
27.2095 | -103.5563
27.2131 | -103.5515
27.2168 | -103.5467
27.2204 | -103.5417
27.224 | -103.5368
27.2277 | -103.5305
27.2313 | -103.5248
27.235 | -103.5215
27.2386 | -103.5174
27.2423 | -103.5135
27.2459 | -103.5093
27.2495 | -103.5037
27.2532 | -103.4981
27.2568 | -103.4929
27.2605 | -103.4878
27.2641 | -103.4819
27.2678 | -103.4758
27.2714 | -103.4705
27.275 | -103.467
27.2787 | -103.4634
27.2823 | -103.4591
27.286 | -103.4535
27.2896 | -103.4458
27.2933 | -103.4374
27.2969 | -103.4302
27.3005 | -103.4238
27.3042 | -103.4167
27.3078 | -103.4102
27.3115 | -103.4028
27.3151 | -103.3946
27.3188 | -103.386
27.3224 | -103.3777
27.326 | -103.3676
27.3297 | -103.3582
27.3333 | -103.3491
27.337 | -103.3395
27.3406 | -103.33
27.3443 | -103.3199
27.3479 | -103.3115
27.3515 | -103.3029
27.3552 | -103.2925
27.3588 | -103.2827
27.3625 | -103.2734
27.3661 | -103.2639
27.3698 | -103.2543
27.3734 | -103.2447
27.377 | -103.2344
27.3807 | -103.2246
27.3843 | -103.2166
27.388 | -103.2057
27.3916 | -103.1957
27.3953 | -103.1862
27.3989 | -103.1763
27.4026 | -103.1686
27.4062 | -103.1641
27.4098 | -103.1588
27.4135 | -103.1514
27.4171 | -103.1434
27.4208 | -103.1358
27.4244 | -103.1284
27.4281 | -103.1218
27.4317 | -103.1151
27.4353 | -103.1092
27.439 | -103.1037
27.4426 | -103.0994
27.4463 | -103.0955
27.4499 | -103.0921
27.4536 | -103.0878
27.4572 | -103.0837
27.4608 | -103.0786
27.4645 | -103.0746
27.4681 | -103.0711
27.4718 | -103.0674
27.4754 | -103.0645
27.4791 | -103.0635
27.4827 | -103.0616
27.4863 | -103.0572
27.49 | -103.0541
27.4936 | -103.0524
27.4973 | -103.0494
27.5009 | -103.0455
27.5046 | -103.0419
27.5082 | -103.0386
27.5118 | -103.0357
27.5155 | -103.0332
27.5191 | -103.0288
27.5228 | -103.0241
27.5264 | -103.0192
27.5301 | -103.0132
27.5337 | -103.0092
27.5373 | -103.0045
27.541 | -102.9974
27.5446 | -102.9922
27.5483 | -102.9873
27.5519 | -102.9806
27.5556 | -102.9738
27.5592 | -102.9674
27.5628 | -102.9608
27.5665 | -102.954
27.5701 | -102.9463
27.5738 | -102.9388
27.5774 | -102.9319
27.5811 | -102.9249
27.5847 | -102.9173
27.5883 | -102.9105
27.592 | -102.9057
27.5956 | -102.8982
27.5993 | -102.8902
27.6029 | -102.8817
27.6066 | -102.8743
27.6102 | -102.8677
27.6138 | -102.8606
27.6175 | -102.8531
27.6211 | -102.8461
27.6248 | -102.8399
27.6284 | -102.8334
27.6321 | -102.8272
27.6357 | -102.8224
27.6393 | -102.8184
27.643 | -102.8145
27.6466 | -102.8106
27.6503 | -102.8054
27.6539 | -102.8004
27.6576 | -102.7966
27.6612 | -102.7928
27.6648 | -102.7896
27.6685 | -102.7879
27.6721 | -102.7883
27.6758 | -102.7888
27.6794 | -102.7897
27.6831 | -102.7902
27.6867 | -102.7918
27.6903 | -102.7928
27.694 | -102.7943
27.6976 | -102.7973
27.7013 | -102.8014
27.7049 | -102.8063
27.7086 | -102.8105
27.7122 | -102.814
27.7158 | -102.8171
27.7195 | -102.8224
27.7231 | -102.8276
27.7268 | -102.8339
27.7304 | -102.8399
27.7341 | -102.847
27.7377 | -102.8546
27.7413 | -102.8621
27.745 | -102.8716
27.7486 | -102.8813
27.7523 | -102.8914
27.7559 | -102.9011
27.7596 | -102.9111
27.7632 | -102.9218
27.7668 | -102.9328
27.7705 | -102.944
27.7741 | -102.954
27.7778 | -102.9645
27.7814 | -102.9759
27.7851 | -102.9866
27.7887 | -102.9963
27.7923 | -103.0052
27.796 | -103.0131
27.7996 | -103.0219
27.8033 | -103.0327
27.8069 | -103.0426
27.8106 | -103.0521
27.8142 | -103.0612
27.8179 | -103.0711
27.8215 | -103.0804
27.8251 | -103.0881
27.8288 | -103.0965
27.8324 | -103.1045
27.8361 | -103.1114
27.8397 | -103.1181
27.8434 | -103.1255
27.847 | -103.1338
27.8506 | -103.143
27.8543 | -103.1486
27.8579 | -103.1559
27.8616 | -103.1642
27.8652 | -103.1721
27.8689 | -103.1789
27.8725 | -103.1854
27.8761 | -103.1921
27.8798 | -103.2
27.8834 | -103.207
27.8871 | -103.2135
27.8907 | -103.2217
27.8944 | -103.2308
27.898 | -103.2403
27.9016 | -103.2485
27.9053 | -103.2559
27.9089 | -103.2633
27.9126 | -103.273
27.9162 | -103.2827
27.9199 | -103.2924
27.9235 | -103.3012
27.9271 | -103.3071
27.9308 | -103.3153
27.9344 | -103.3253
27.9381 | -103.3329
27.9417 | -103.3406
27.9454 | -103.3497
27.949 | -103.3579
27.9526 | -103.3646
27.9563 | -103.3715
27.9599 | -103.3795
27.9636 | -103.3889
27.9672 | -103.3982
27.9709 | -103.4076
27.9745 | -103.4168
27.9781 | -103.4265
27.9818 | -103.436
27.9854 | -103.4452
27.9891 | -103.4552
27.9927 | -103.466
27.9964 | -103.4759
28 | -103.4846
28.0036 | -103.4915