## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 248 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 266
Hood | 02-92B/P | 192
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1011.5 | hPa |  | PRC PRO 01
Humidity | 56.4 | %/Hr
Internal temperature of calibration bench | 23.41 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1019 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.1 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.02 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.03 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3072 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 13/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 19/06/2018
Technical Manager | Anthony HUE | Approval |  | 28/06/2018
CEO | Robert GLEMOT | Approval |  | 28/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3922 | 57.5307
0.012209 | -35.9532 | 38.353
0.018313 | -34.8961 | 28.2421
0.024417 | -34.4337 | 21.9568
0.030521 | -34.3197 | 18.168
0.036626 | -34.1754 | 15.1087
0.04273 | -34.0799 | 12.7462
0.048834 | -34.0343 | 11.2129
0.054938 | -33.9718 | 10.0715
0.061043 | -33.9129 | 8.9907
0.067147 | -33.9018 | 8.1688
0.073251 | -33.9027 | 7.4086
0.079355 | -33.921 | 6.7484
0.08546 | -33.9347 | 6.1591
0.091564 | -33.9296 | 5.7218
0.097668 | -33.912 | 5.4723
0.10377 | -33.8916 | 5.2197
0.10988 | -33.8755 | 4.9388
0.11598 | -33.867 | 4.6451
0.12209 | -33.8691 | 4.3644
0.12819 | -33.879 | 4.0867
0.13429 | -33.8899 | 3.8039
0.1404 | -33.8935 | 3.5297
0.1465 | -33.8957 | 3.2761
0.15261 | -33.8935 | 3.0594
0.15871 | -33.8875 | 2.8982
0.16482 | -33.8799 | 2.7941
0.17092 | -33.8775 | 2.751
0.17702 | -33.8754 | 2.6763
0.18313 | -33.8745 | 2.5964
0.18923 | -33.874 | 2.4878
0.19534 | -33.8727 | 2.3627
0.20144 | -33.8712 | 2.2237
0.20754 | -33.868 | 2.1066
0.21365 | -33.8668 | 1.995
0.21975 | -33.8662 | 1.8993
0.22586 | -33.8672 | 1.8133
0.23196 | -33.8697 | 1.7408
0.23807 | -33.874 | 1.6769
0.24417 | -33.8783 | 1.6159
0.25027 | -33.8824 | 1.5519
0.25638 | -33.885 | 1.4887
0.26248 | -33.8869 | 1.4215
0.26859 | -33.8879 | 1.3562
0.27469 | -33.8883 | 1.2861
0.2808 | -33.8893 | 1.215
0.2869 | -33.8901 | 1.1459
0.293 | -33.8911 | 1.0808
0.29911 | -33.892 | 1.0227
0.30521 | -33.893 | 0.96538
0.31132 | -33.8937 | 0.92715
0.31742 | -33.8949 | 0.88508
0.32353 | -33.8953 | 0.85038
0.32963 | -33.8962 | 0.82124
0.33573 | -33.8973 | 0.80028
0.34184 | -33.8987 | 0.77821
0.34794 | -33.9009 | 0.75383
0.35405 | -33.9031 | 0.72923
0.36015 | -33.9056 | 0.70339
0.36626 | -33.9077 | 0.67555
0.37236 | -33.9099 | 0.63886
0.37846 | -33.912 | 0.59496
0.38457 | -33.9129 | 0.5387
0.39067 | -33.914 | 0.49259
0.39678 | -33.9147 | 0.44148
0.40288 | -33.9151 | 0.38829
0.40899 | -33.915 | 0.33707
0.41509 | -33.9144 | 0.28712
0.42119 | -33.9144 | 0.24463
0.4273 | -33.9143 | 0.19919
0.4334 | -33.9143 | 0.15384
0.43951 | -33.9146 | 0.11107
0.44561 | -33.9155 | 0.07187
0.45172 | -33.9171 | 0.040378
0.45782 | -33.9179 | 0.0037909
0.46392 | -33.9192 | -0.032023
0.47003 | -33.9203 | -0.06414
0.47613 | -33.9212 | -0.097932
0.48224 | -33.9218 | -0.13239
0.48834 | -33.921 | -0.16699
0.49445 | -33.9209 | -0.20191
0.50055 | -33.9204 | -0.23386
0.50665 | -33.9198 | -0.26787
0.51276 | -33.9189 | -0.30175
0.51886 | -33.918 | -0.33551
0.52497 | -33.9168 | -0.3598
0.53107 | -33.9161 | -0.38927
0.53717 | -33.9155 | -0.41748
0.54328 | -33.9147 | -0.44046
0.54938 | -33.9146 | -0.4631
0.55549 | -33.9147 | -0.48405
0.56159 | -33.9158 | -0.49582
0.5677 | -33.9162 | -0.51053
0.5738 | -33.9172 | -0.52698
0.5799 | -33.9181 | -0.54201
0.58601 | -33.9192 | -0.55231
0.59211 | -33.9208 | -0.5587
0.59822 | -33.9222 | -0.56828
0.60432 | -33.9237 | -0.58043
0.61043 | -33.9249 | -0.58856
0.61653 | -33.9262 | -0.59887
0.62263 | -33.9277 | -0.60824
0.62874 | -33.9283 | -0.61887
0.63484 | -33.9292 | -0.62821
0.64095 | -33.9297 | -0.6384
0.64705 | -33.9301 | -0.64825
0.65316 | -33.9303 | -0.66065
0.65926 | -33.9304 | -0.67636
0.66536 | -33.9293 | -0.69381
0.67147 | -33.9289 | -0.70709
0.67757 | -33.9286 | -0.72147
0.68368 | -33.9279 | -0.74031
0.68978 | -33.9272 | -0.75754
0.69589 | -33.9263 | -0.77655
0.70199 | -33.9254 | -0.79915
0.70809 | -33.9246 | -0.81826
0.7142 | -33.9239 | -0.83862
0.7203 | -33.9229 | -0.85783
0.72641 | -33.9223 | -0.88011
0.73251 | -33.9216 | -0.90009
0.73862 | -33.9219 | -0.92167
0.74472 | -33.9216 | -0.94158
0.75082 | -33.9214 | -0.96349
0.75693 | -33.9217 | -0.98515
0.76303 | -33.9222 | -1.0069
0.76914 | -33.924 | -1.0271
0.77524 | -33.925 | -1.0466
0.78135 | -33.9262 | -1.0665
0.78745 | -33.9276 | -1.0862
0.79355 | -33.9288 | -1.1058
0.79966 | -33.9305 | -1.1251
0.80576 | -33.9326 | -1.1445
0.81187 | -33.9342 | -1.1659
0.81797 | -33.9361 | -1.1869
0.82408 | -33.9377 | -1.2071
0.83018 | -33.9394 | -1.2274
0.83628 | -33.9412 | -1.2457
0.84239 | -33.9423 | -1.266
0.84849 | -33.9437 | -1.2875
0.8546 | -33.9451 | -1.3071
0.8607 | -33.9463 | -1.3276
0.86681 | -33.9474 | -1.3464
0.87291 | -33.9483 | -1.3651
0.87901 | -33.9488 | -1.3831
0.88512 | -33.9495 | -1.4012
0.89122 | -33.9502 | -1.4193
0.89733 | -33.9506 | -1.435
0.90343 | -33.9511 | -1.4525
0.90953 | -33.9512 | -1.4669
0.91564 | -33.9513 | -1.4828
0.92174 | -33.9515 | -1.4978
0.92785 | -33.9516 | -1.5124
0.93395 | -33.9517 | -1.5263
0.94006 | -33.9519 | -1.5392
0.94616 | -33.9517 | -1.5511
0.95226 | -33.9518 | -1.5639
0.95837 | -33.9519 | -1.5777
0.96447 | -33.9522 | -1.5905
0.97058 | -33.952 | -1.6039
0.97668 | -33.9519 | -1.6172
0.98279 | -33.9516 | -1.6299
0.98889 | -33.9515 | -1.6406
0.99499 | -33.9515 | -1.654
1.0011 | -33.9514 | -1.6647
1.0072 | -33.9511 | -1.6763
1.0133 | -33.9508 | -1.687
1.0194 | -33.9504 | -1.6999
1.0255 | -33.9501 | -1.7135
1.0316 | -33.9496 | -1.7254
1.0377 | -33.9492 | -1.7389
1.0438 | -33.9487 | -1.7515
1.0499 | -33.948 | -1.7666
1.056 | -33.9473 | -1.7807
1.0621 | -33.9467 | -1.7948
1.0682 | -33.9463 | -1.8071
1.0743 | -33.9456 | -1.8224
1.0805 | -33.9452 | -1.8366
1.0866 | -33.9443 | -1.8515
1.0927 | -33.9435 | -1.8657
1.0988 | -33.9429 | -1.8804
1.1049 | -33.9421 | -1.8945
1.111 | -33.9413 | -1.9096
1.1171 | -33.9407 | -1.924
1.1232 | -33.9402 | -1.9389
1.1293 | -33.9398 | -1.9528
1.1354 | -33.9395 | -1.9676
1.1415 | -33.9392 | -1.9829
1.1476 | -33.9386 | -1.9982
1.1537 | -33.9381 | -2.0127
1.1598 | -33.9382 | -2.0272
1.1659 | -33.938 | -2.0404
1.172 | -33.9378 | -2.0542
1.1781 | -33.9375 | -2.068
1.1842 | -33.9374 | -2.082
1.1903 | -33.9378 | -2.093
1.1964 | -33.9377 | -2.1063
1.2025 | -33.9378 | -2.1201
1.2086 | -33.9381 | -2.1325
1.2147 | -33.9382 | -2.1449
1.2209 | -33.9385 | -2.1561
1.227 | -33.9393 | -2.1668
1.2331 | -33.9397 | -2.1787
1.2392 | -33.94 | -2.1905
1.2453 | -33.9405 | -2.2016
1.2514 | -33.9409 | -2.2115
1.2575 | -33.9415 | -2.2224
1.2636 | -33.9423 | -2.2316
1.2697 | -33.9429 | -2.2409
1.2758 | -33.9436 | -2.2521
1.2819 | -33.9443 | -2.2621
1.288 | -33.945 | -2.2729
1.2941 | -33.9458 | -2.2815
1.3002 | -33.9465 | -2.2898
1.3063 | -33.947 | -2.2989
1.3124 | -33.9477 | -2.3081
1.3185 | -33.9484 | -2.3172
1.3246 | -33.9491 | -2.3265
1.3307 | -33.9499 | -2.3352
1.3368 | -33.9507 | -2.344
1.3429 | -33.9513 | -2.3531
1.349 | -33.9518 | -2.3625
1.3551 | -33.9525 | -2.372
1.3613 | -33.953 | -2.3804
1.3674 | -33.9536 | -2.3905
1.3735 | -33.9541 | -2.4003
1.3796 | -33.9545 | -2.4095
1.3857 | -33.9549 | -2.4188
1.3918 | -33.9554 | -2.429
1.3979 | -33.9559 | -2.4382
1.404 | -33.9564 | -2.4495
1.4101 | -33.9568 | -2.4599
1.4162 | -33.9572 | -2.4705
1.4223 | -33.9576 | -2.4818
1.4284 | -33.9581 | -2.4918
1.4345 | -33.9583 | -2.5033
1.4406 | -33.9586 | -2.5166
1.4467 | -33.9591 | -2.5281
1.4528 | -33.9595 | -2.5394
1.4589 | -33.96 | -2.5511
1.465 | -33.9602 | -2.5619
1.4711 | -33.9607 | -2.5732
1.4772 | -33.9611 | -2.5863
1.4833 | -33.9616 | -2.5992
1.4894 | -33.9621 | -2.6112
1.4955 | -33.9626 | -2.623
1.5016 | -33.9631 | -2.6354
1.5078 | -33.9635 | -2.649
1.5139 | -33.9639 | -2.6618
1.52 | -33.9643 | -2.6743
1.5261 | -33.9646 | -2.6867
1.5322 | -33.9649 | -2.7005
1.5383 | -33.9653 | -2.7129
1.5444 | -33.9658 | -2.7254
1.5505 | -33.9662 | -2.7386
1.5566 | -33.9667 | -2.7514
1.5627 | -33.9672 | -2.765
1.5688 | -33.9678 | -2.7784
1.5749 | -33.9681 | -2.7908
1.581 | -33.9684 | -2.804
1.5871 | -33.9688 | -2.8175
1.5932 | -33.9691 | -2.8302
1.5993 | -33.9695 | -2.8442
1.6054 | -33.9699 | -2.8572
1.6115 | -33.9703 | -2.8712
1.6176 | -33.9706 | -2.8836
1.6237 | -33.9708 | -2.897
1.6298 | -33.971 | -2.9094
1.6359 | -33.9714 | -2.922
1.642 | -33.9718 | -2.9338
1.6482 | -33.972 | -2.9466
1.6543 | -33.9723 | -2.9601
1.6604 | -33.9726 | -2.972
1.6665 | -33.9727 | -2.9842
1.6726 | -33.9729 | -2.9966
1.6787 | -33.9733 | -3.009
1.6848 | -33.9732 | -3.0214
1.6909 | -33.9734 | -3.0339
1.697 | -33.9737 | -3.046
1.7031 | -33.9737 | -3.0584
1.7092 | -33.9737 | -3.0704
1.7153 | -33.9738 | -3.0822
1.7214 | -33.9736 | -3.0943
1.7275 | -33.9736 | -3.1068
1.7336 | -33.9735 | -3.1195
1.7397 | -33.9737 | -3.1317
1.7458 | -33.9737 | -3.1437
1.7519 | -33.9737 | -3.1558
1.758 | -33.9737 | -3.1676
1.7641 | -33.9739 | -3.1791
1.7702 | -33.9738 | -3.1919
1.7763 | -33.9738 | -3.2033
1.7824 | -33.9738 | -3.2156
1.7885 | -33.9737 | -3.2279
1.7947 | -33.9739 | -3.2398
1.8008 | -33.9739 | -3.2509
1.8069 | -33.9739 | -3.2624
1.813 | -33.974 | -3.2744
1.8191 | -33.9739 | -3.287
1.8252 | -33.974 | -3.2996
1.8313 | -33.9741 | -3.3117
1.8374 | -33.9743 | -3.3235
1.8435 | -33.9743 | -3.3354
1.8496 | -33.9744 | -3.3483
1.8557 | -33.9746 | -3.3609
1.8618 | -33.9748 | -3.3731
1.8679 | -33.9749 | -3.385
1.874 | -33.975 | -3.3969
1.8801 | -33.9752 | -3.4093
1.8862 | -33.9754 | -3.4212
1.8923 | -33.9755 | -3.4333
1.8984 | -33.9756 | -3.4453
1.9045 | -33.9757 | -3.4569
1.9106 | -33.9758 | -3.469
1.9167 | -33.9758 | -3.4812
1.9228 | -33.9759 | -3.493
1.9289 | -33.976 | -3.5048
1.9351 | -33.9762 | -3.5157
1.9412 | -33.9764 | -3.5269
1.9473 | -33.9767 | -3.539
1.9534 | -33.9769 | -3.5504
1.9595 | -33.9771 | -3.5622
1.9656 | -33.9773 | -3.5745
1.9717 | -33.9776 | -3.5862
1.9778 | -33.9779 | -3.5976
1.9839 | -33.9781 | -3.6091
1.99 | -33.9783 | -3.6202
1.9961 | -33.9784 | -3.6323
2.0022 | -33.9786 | -3.6437
2.0083 | -33.9787 | -3.655
2.0144 | -33.9789 | -3.6664
2.0205 | -33.9792 | -3.6781
2.0266 | -33.9795 | -3.6895
2.0327 | -33.9797 | -3.7011
2.0388 | -33.9801 | -3.7127
2.0449 | -33.9803 | -3.7242
2.051 | -33.9806 | -3.7356
2.0571 | -33.9809 | -3.7469
2.0632 | -33.9811 | -3.7584
2.0693 | -33.9813 | -3.7706
2.0754 | -33.9815 | -3.7819
2.0816 | -33.9819 | -3.7941
2.0877 | -33.982 | -3.8056
2.0938 | -33.9824 | -3.8177
2.0999 | -33.9826 | -3.8294
2.106 | -33.9828 | -3.841
2.1121 | -33.983 | -3.8529
2.1182 | -33.9832 | -3.8638
2.1243 | -33.9835 | -3.8745
2.1304 | -33.9837 | -3.8857
2.1365 | -33.984 | -3.8975
2.1426 | -33.9842 | -3.9094
2.1487 | -33.9845 | -3.9207
2.1548 | -33.9849 | -3.9325
2.1609 | -33.9853 | -3.9439
2.167 | -33.9857 | -3.9552
2.1731 | -33.986 | -3.9671
2.1792 | -33.9862 | -3.9782
2.1853 | -33.9866 | -3.9901
2.1914 | -33.9869 | -4.0015
2.1975 | -33.9873 | -4.013
2.2036 | -33.9875 | -4.0251
2.2097 | -33.9877 | -4.0366
2.2158 | -33.9881 | -4.0472
2.222 | -33.9884 | -4.0577
2.2281 | -33.9886 | -4.0692
2.2342 | -33.9889 | -4.0805
2.2403 | -33.9891 | -4.0911
2.2464 | -33.9894 | -4.1019
2.2525 | -33.9897 | -4.1132
2.2586 | -33.99 | -4.1238
2.2647 | -33.9903 | -4.1343
2.2708 | -33.9906 | -4.1455
2.2769 | -33.991 | -4.1562
2.283 | -33.9912 | -4.1673
2.2891 | -33.9914 | -4.1782
2.2952 | -33.9916 | -4.1888
2.3013 | -33.9918 | -4.1996
2.3074 | -33.9921 | -4.2101
2.3135 | -33.9923 | -4.2215
2.3196 | -33.9925 | -4.2317
2.3257 | -33.9927 | -4.2424
2.3318 | -33.9929 | -4.2535
2.3379 | -33.9931 | -4.2642
2.344 | -33.9932 | -4.2756
2.3501 | -33.9934 | -4.2862
2.3562 | -33.9936 | -4.2966
2.3623 | -33.9937 | -4.307
2.3685 | -33.9939 | -4.3175
2.3746 | -33.994 | -4.3287
2.3807 | -33.9942 | -4.3392
2.3868 | -33.9944 | -4.3499
2.3929 | -33.9946 | -4.3605
2.399 | -33.9947 | -4.371
2.4051 | -33.995 | -4.3815
2.4112 | -33.9952 | -4.3921
2.4173 | -33.9953 | -4.4027
2.4234 | -33.9954 | -4.4135
2.4295 | -33.9956 | -4.4243
2.4356 | -33.9957 | -4.4351
2.4417 | -33.9959 | -4.4462
2.4478 | -33.996 | -4.4571
2.4539 | -33.9962 | -4.4673
2.46 | -33.9963 | -4.4786
2.4661 | -33.9964 | -4.4897
2.4722 | -33.9966 | -4.4998
2.4783 | -33.9967 | -4.5102
2.4844 | -33.9967 | -4.5212
2.4905 | -33.9968 | -4.5321
2.4966 | -33.997 | -4.5429
2.5027 | -33.9971 | -4.5533
2.5089 | -33.9972 | -4.5641
2.515 | -33.9972 | -4.5749
2.5211 | -33.9973 | -4.5856
2.5272 | -33.9974 | -4.5964
2.5333 | -33.9975 | -4.6073
2.5394 | -33.9977 | -4.618
2.5455 | -33.9978 | -4.6293
2.5516 | -33.9979 | -4.6396
2.5577 | -33.9981 | -4.651
2.5638 | -33.9983 | -4.6626
2.5699 | -33.9983 | -4.674
2.576 | -33.9984 | -4.6855
2.5821 | -33.9985 | -4.6966
2.5882 | -33.9986 | -4.7077
2.5943 | -33.9987 | -4.7185
2.6004 | -33.9989 | -4.7294
2.6065 | -33.999 | -4.7398
2.6126 | -33.9991 | -4.7502
2.6187 | -33.9992 | -4.7606
2.6248 | -33.9993 | -4.7713
2.6309 | -33.9995 | -4.7827
2.637 | -33.9997 | -4.7942
2.6431 | -33.9998 | -4.8046
2.6492 | -33.9999 | -4.8159
2.6554 | -34.0001 | -4.8268
2.6615 | -34.0002 | -4.8376
2.6676 | -34.0004 | -4.849
2.6737 | -34.0007 | -4.8601
2.6798 | -34.001 | -4.8711
2.6859 | -34.0011 | -4.8823
2.692 | -34.0013 | -4.8931
2.6981 | -34.0015 | -4.9042
2.7042 | -34.0016 | -4.9153
2.7103 | -34.0018 | -4.9259
2.7164 | -34.0021 | -4.9368
2.7225 | -34.0023 | -4.9481
2.7286 | -34.0026 | -4.9591
2.7347 | -34.0027 | -4.9696
2.7408 | -34.003 | -4.9797
2.7469 | -34.0032 | -4.9909
2.753 | -34.0035 | -5.0015
2.7591 | -34.0038 | -5.0128
2.7652 | -34.004 | -5.0242
2.7713 | -34.0041 | -5.0349
2.7774 | -34.0044 | -5.0458
2.7835 | -34.0045 | -5.0567
2.7896 | -34.0047 | -5.0673
2.7958 | -34.005 | -5.078
2.8019 | -34.0053 | -5.0884
2.808 | -34.0055 | -5.0992
2.8141 | -34.0058 | -5.1103
2.8202 | -34.0061 | -5.1211
2.8263 | -34.0064 | -5.132
2.8324 | -34.0067 | -5.1433
2.8385 | -34.0071 | -5.1542
2.8446 | -34.0074 | -5.1649
2.8507 | -34.0078 | -5.1758
2.8568 | -34.0081 | -5.1862
2.8629 | -34.0084 | -5.1969
2.869 | -34.0087 | -5.208
2.8751 | -34.009 | -5.2187
2.8812 | -34.0094 | -5.23
2.8873 | -34.0097 | -5.2409
2.8934 | -34.01 | -5.2519
2.8995 | -34.0103 | -5.263
2.9056 | -34.0106 | -5.2741
2.9117 | -34.0109 | -5.2851
2.9178 | -34.0113 | -5.2959
2.9239 | -34.0117 | -5.3074
2.93 | -34.0121 | -5.3188
2.9361 | -34.0123 | -5.3296
2.9423 | -34.0127 | -5.3403
2.9484 | -34.0131 | -5.3514
2.9545 | -34.0135 | -5.3624
2.9606 | -34.0139 | -5.3738
2.9667 | -34.0142 | -5.385
2.9728 | -34.0145 | -5.3963
2.9789 | -34.015 | -5.4073
2.985 | -34.0154 | -5.4187
2.9911 | -34.0157 | -5.43
2.9972 | -34.0162 | -5.441
3.0033 | -34.0166 | -5.4525
3.0094 | -34.0169 | -5.4636
3.0155 | -34.0174 | -5.4746
3.0216 | -34.0177 | -5.4858
3.0277 | -34.018 | -5.4974
3.0338 | -34.0184 | -5.5088
3.0399 | -34.0188 | -5.5201
3.046 | -34.0191 | -5.5317
3.0521 | -34.0195 | -5.5428
3.0582 | -34.0198 | -5.5548
3.0643 | -34.0201 | -5.5663
3.0704 | -34.0204 | -5.5781
3.0765 | -34.0207 | -5.5892
3.0827 | -34.021 | -5.6004
3.0888 | -34.0213 | -5.6117
3.0949 | -34.0217 | -5.623
3.101 | -34.0221 | -5.6345
3.1071 | -34.0224 | -5.6462
3.1132 | -34.0228 | -5.658
3.1193 | -34.0231 | -5.6694
3.1254 | -34.0235 | -5.6807
3.1315 | -34.0238 | -5.6926
3.1376 | -34.0241 | -5.704
3.1437 | -34.0245 | -5.7155
3.1498 | -34.0248 | -5.7275
3.1559 | -34.0251 | -5.7389
3.162 | -34.0254 | -5.751
3.1681 | -34.0256 | -5.7629
3.1742 | -34.0259 | -5.7744
3.1803 | -34.0263 | -5.786
3.1864 | -34.0266 | -5.7973
3.1925 | -34.0269 | -5.8088
3.1986 | -34.0272 | -5.8207
3.2047 | -34.0274 | -5.8327
3.2108 | -34.0277 | -5.8438
3.2169 | -34.028 | -5.8555
3.223 | -34.0283 | -5.8672
3.2292 | -34.0285 | -5.8789
3.2353 | -34.0287 | -5.8897
3.2414 | -34.0289 | -5.9011
3.2475 | -34.0291 | -5.9125
3.2536 | -34.0295 | -5.9239
3.2597 | -34.0297 | -5.9349
3.2658 | -34.0299 | -5.9458
3.2719 | -34.0301 | -5.9567
3.278 | -34.0304 | -5.9683
3.2841 | -34.0306 | -5.9794
3.2902 | -34.0309 | -5.9908
3.2963 | -34.0311 | -6.0022
3.3024 | -34.0312 | -6.0133
3.3085 | -34.0315 | -6.0248
3.3146 | -34.0317 | -6.0359
3.3207 | -34.0319 | -6.0473
3.3268 | -34.0321 | -6.0587
3.3329 | -34.0323 | -6.0695
3.339 | -34.0325 | -6.0803
3.3451 | -34.0327 | -6.0915
3.3512 | -34.0329 | -6.1022
3.3573 | -34.0331 | -6.1134
3.3634 | -34.0333 | -6.1245
3.3696 | -34.0335 | -6.1358
3.3757 | -34.0336 | -6.1465
3.3818 | -34.0339 | -6.1573
3.3879 | -34.034 | -6.1683
3.394 | -34.0342 | -6.1796
3.4001 | -34.0344 | -6.1904
3.4062 | -34.0345 | -6.2013
3.4123 | -34.0346 | -6.2119
3.4184 | -34.0348 | -6.2231
3.4245 | -34.035 | -6.2339
3.4306 | -34.0351 | -6.2445
3.4367 | -34.0353 | -6.255
3.4428 | -34.0355 | -6.2654
3.4489 | -34.0356 | -6.2763
3.455 | -34.0358 | -6.2871
3.4611 | -34.036 | -6.2983
3.4672 | -34.0361 | -6.3093
3.4733 | -34.0363 | -6.3203
3.4794 | -34.0364 | -6.3311
3.4855 | -34.0366 | -6.3414
3.4916 | -34.0368 | -6.3518
3.4977 | -34.0369 | -6.363
3.5038 | -34.0371 | -6.3739
3.5099 | -34.0372 | -6.3841
3.5161 | -34.0374 | -6.3945
3.5222 | -34.0375 | -6.4048
3.5283 | -34.0376 | -6.4156
3.5344 | -34.0377 | -6.4263
3.5405 | -34.0379 | -6.4368
3.5466 | -34.0381 | -6.4477
3.5527 | -34.0382 | -6.4585
3.5588 | -34.0383 | -6.4691
3.5649 | -34.0384 | -6.4797
3.571 | -34.0385 | -6.4903
3.5771 | -34.0386 | -6.5014
3.5832 | -34.0388 | -6.5113
3.5893 | -34.039 | -6.5219
3.5954 | -34.0392 | -6.5327
3.6015 | -34.0393 | -6.5436
3.6076 | -34.0395 | -6.5541
3.6137 | -34.0397 | -6.5642
3.6198 | -34.0399 | -6.5745
3.6259 | -34.0401 | -6.5853
3.632 | -34.0403 | -6.5959
3.6381 | -34.0404 | -6.6067
3.6442 | -34.0406 | -6.6172
3.6503 | -34.0408 | -6.6274
3.6565 | -34.041 | -6.6382
3.6626 | -34.0411 | -6.6488
3.6687 | -34.0413 | -6.6595
3.6748 | -34.0416 | -6.6699
3.6809 | -34.0418 | -6.6805
3.687 | -34.042 | -6.6914
3.6931 | -34.0423 | -6.7017
3.6992 | -34.0425 | -6.7123
3.7053 | -34.0427 | -6.7228
3.7114 | -34.0428 | -6.7336
3.7175 | -34.043 | -6.7443
3.7236 | -34.0433 | -6.7543
3.7297 | -34.0435 | -6.7651
3.7358 | -34.0437 | -6.776
3.7419 | -34.0439 | -6.7868
3.748 | -34.0441 | -6.7974
3.7541 | -34.0443 | -6.808
3.7602 | -34.0446 | -6.8185
3.7663 | -34.0448 | -6.8289
3.7724 | -34.0451 | -6.8396
3.7785 | -34.0453 | -6.85
3.7846 | -34.0455 | -6.8604
3.7907 | -34.0458 | -6.8714
3.7969 | -34.0461 | -6.8821
3.803 | -34.0463 | -6.8924
3.8091 | -34.0465 | -6.903
3.8152 | -34.0468 | -6.9135
3.8213 | -34.047 | -6.9243
3.8274 | -34.0473 | -6.9346
3.8335 | -34.0476 | -6.9454
3.8396 | -34.0479 | -6.956
3.8457 | -34.0482 | -6.9668
3.8518 | -34.0486 | -6.9774
3.8579 | -34.0489 | -6.988
3.864 | -34.0492 | -6.9988
3.8701 | -34.0495 | -7.0098
3.8762 | -34.0497 | -7.0203
3.8823 | -34.05 | -7.0311
3.8884 | -34.0503 | -7.0418
3.8945 | -34.0506 | -7.0528
3.9006 | -34.0509 | -7.0638
3.9067 | -34.0512 | -7.0749
3.9128 | -34.0515 | -7.0857
3.9189 | -34.0518 | -7.0962
3.925 | -34.052 | -7.107
3.9311 | -34.0523 | -7.1172
3.9372 | -34.0526 | -7.1281
3.9434 | -34.0529 | -7.1389
3.9495 | -34.0532 | -7.1498
3.9556 | -34.0536 | -7.1603
3.9617 | -34.0539 | -7.1707
3.9678 | -34.0542 | -7.1811
3.9739 | -34.0544 | -7.1918
3.98 | -34.0547 | -7.2027
3.9861 | -34.055 | -7.2133
3.9922 | -34.0552 | -7.224
3.9983 | -34.0555 | -7.2344
4.0044 | -34.0558 | -7.2447
4.0105 | -34.0561 | -7.2554
4.0166 | -34.0564 | -7.2661
4.0227 | -34.0567 | -7.2768
4.0288 | -34.057 | -7.288
4.0349 | -34.0572 | -7.2984
4.041 | -34.0575 | -7.3091
4.0471 | -34.0577 | -7.3199
4.0532 | -34.058 | -7.3302
4.0593 | -34.0583 | -7.3411
4.0654 | -34.0586 | -7.3515
4.0715 | -34.0589 | -7.3627
4.0776 | -34.0592 | -7.3733
4.0838 | -34.0595 | -7.3843
4.0899 | -34.0598 | -7.3948
4.096 | -34.0601 | -7.4054
4.1021 | -34.0603 | -7.4159
4.1082 | -34.0606 | -7.4264
4.1143 | -34.0608 | -7.4367
4.1204 | -34.0611 | -7.4471
4.1265 | -34.0614 | -7.4575
4.1326 | -34.0617 | -7.4678
4.1387 | -34.062 | -7.4782
4.1448 | -34.0623 | -7.4889
4.1509 | -34.0625 | -7.4996
4.157 | -34.0628 | -7.5106
4.1631 | -34.063 | -7.5214
4.1692 | -34.0633 | -7.5319
4.1753 | -34.0636 | -7.5425
4.1814 | -34.0638 | -7.5527
4.1875 | -34.0641 | -7.5635
4.1936 | -34.0644 | -7.5742
4.1997 | -34.0646 | -7.5845
4.2058 | -34.0648 | -7.5953
4.2119 | -34.0651 | -7.6058
4.218 | -34.0653 | -7.6161
4.2241 | -34.0655 | -7.6264
4.2303 | -34.0658 | -7.6368
4.2364 | -34.0661 | -7.6473
4.2425 | -34.0663 | -7.6579
4.2486 | -34.0665 | -7.6686
4.2547 | -34.0666 | -7.6789
4.2608 | -34.0669 | -7.6895
4.2669 | -34.0671 | -7.6996
4.273 | -34.0673 | -7.7103
4.2791 | -34.0676 | -7.7208
4.2852 | -34.0678 | -7.7315
4.2913 | -34.0679 | -7.7416
4.2974 | -34.0682 | -7.7524
4.3035 | -34.0684 | -7.7632
4.3096 | -34.0686 | -7.7739
4.3157 | -34.0689 | -7.7842
4.3218 | -34.0691 | -7.7946
4.3279 | -34.0693 | -7.8051
4.334 | -34.0695 | -7.8155
4.3401 | -34.0697 | -7.8255
4.3462 | -34.07 | -7.836
4.3523 | -34.0702 | -7.8461
4.3584 | -34.0704 | -7.8561
4.3645 | -34.0706 | -7.8664
4.3707 | -34.0708 | -7.8765
4.3768 | -34.071 | -7.8869
4.3829 | -34.0713 | -7.8976
4.389 | -34.0715 | -7.908
4.3951 | -34.0717 | -7.9182
4.4012 | -34.0719 | -7.9289
4.4073 | -34.072 | -7.9391
4.4134 | -34.0723 | -7.9494
4.4195 | -34.0725 | -7.9597
4.4256 | -34.0728 | -7.9698
4.4317 | -34.073 | -7.9801
4.4378 | -34.0732 | -7.9901
4.4439 | -34.0734 | -8.0002
4.45 | -34.0735 | -8.0105
4.4561 | -34.0737 | -8.0208
4.4622 | -34.0739 | -8.0311
4.4683 | -34.0741 | -8.0414
4.4744 | -34.0743 | -8.052
4.4805 | -34.0745 | -8.0619
4.4866 | -34.0747 | -8.0722
4.4927 | -34.075 | -8.0827
4.4988 | -34.0752 | -8.093
4.5049 | -34.0754 | -8.1029
4.511 | -34.0756 | -8.1129
4.5172 | -34.0758 | -8.1231
4.5233 | -34.076 | -8.1337
4.5294 | -34.0762 | -8.1438
4.5355 | -34.0764 | -8.1541
4.5416 | -34.0766 | -8.1643
4.5477 | -34.0769 | -8.1745
4.5538 | -34.077 | -8.1843
4.5599 | -34.0773 | -8.1945
4.566 | -34.0775 | -8.2047
4.5721 | -34.0777 | -8.2149
4.5782 | -34.0779 | -8.2254
4.5843 | -34.0781 | -8.2357
4.5904 | -34.0784 | -8.246
4.5965 | -34.0786 | -8.2561
4.6026 | -34.0788 | -8.2668
4.6087 | -34.079 | -8.2768
4.6148 | -34.0792 | -8.2874
4.6209 | -34.0795 | -8.2975
4.627 | -34.0797 | -8.308
4.6331 | -34.0798 | -8.3181
4.6392 | -34.0801 | -8.3284
4.6453 | -34.0803 | -8.3386
4.6514 | -34.0805 | -8.3488
4.6576 | -34.0807 | -8.3594
4.6637 | -34.0809 | -8.3693
4.6698 | -34.0811 | -8.3793
4.6759 | -34.0813 | -8.3899
4.682 | -34.0815 | -8.4
4.6881 | -34.0816 | -8.4102
4.6942 | -34.0818 | -8.4204
4.7003 | -34.082 | -8.4307
4.7064 | -34.0823 | -8.4407
4.7125 | -34.0825 | -8.4511
4.7186 | -34.0827 | -8.4611
4.7247 | -34.0829 | -8.4714
4.7308 | -34.0832 | -8.4815
4.7369 | -34.0834 | -8.4918
4.743 | -34.0836 | -8.5024
4.7491 | -34.0838 | -8.5128
4.7552 | -34.084 | -8.5232
4.7613 | -34.0842 | -8.5335
4.7674 | -34.0844 | -8.5436
4.7735 | -34.0846 | -8.5543
4.7796 | -34.0848 | -8.5648
4.7857 | -34.0851 | -8.5751
4.7918 | -34.0853 | -8.5852
4.7979 | -34.0855 | -8.5954
4.8041 | -34.0857 | -8.6056
4.8102 | -34.086 | -8.6156
4.8163 | -34.0862 | -8.6258
4.8224 | -34.0864 | -8.6361
4.8285 | -34.0866 | -8.646
4.8346 | -34.0869 | -8.6561
4.8407 | -34.0871 | -8.6664
4.8468 | -34.0873 | -8.6769
4.8529 | -34.0876 | -8.6874
4.859 | -34.0878 | -8.6979
4.8651 | -34.088 | -8.7079
4.8712 | -34.0882 | -8.718
4.8773 | -34.0884 | -8.7283
4.8834 | -34.0887 | -8.7386
4.8895 | -34.0889 | -8.7486
4.8956 | -34.0891 | -8.7588
4.9017 | -34.0893 | -8.7689
4.9078 | -34.0895 | -8.7794
4.9139 | -34.0897 | -8.7893
4.92 | -34.09 | -8.7992
4.9261 | -34.0903 | -8.8091
4.9322 | -34.0905 | -8.8194
4.9383 | -34.0907 | -8.8296
4.9445 | -34.091 | -8.8399
4.9506 | -34.0912 | -8.85
4.9567 | -34.0914 | -8.8599
4.9628 | -34.0916 | -8.8701
4.9689 | -34.0919 | -8.8807
4.975 | -34.0921 | -8.8912
4.9811 | -34.0923 | -8.9011
4.9872 | -34.0926 | -8.9113
4.9933 | -34.0929 | -8.9217
4.9994 | -34.0932 | -8.9323
5.0055 | -34.0934 | -8.9425
5.0116 | -34.0937 | -8.9532
5.0177 | -34.094 | -8.9637
5.0238 | -34.0942 | -8.9742
5.0299 | -34.0945 | -8.9848
5.036 | -34.0948 | -8.9951
5.0421 | -34.095 | -9.0053
5.0482 | -34.0953 | -9.0156
5.0543 | -34.0955 | -9.0256
5.0604 | -34.0958 | -9.0358
5.0665 | -34.0961 | -9.0458
5.0726 | -34.0963 | -9.0564
5.0787 | -34.0966 | -9.0668
5.0848 | -34.0969 | -9.0772
5.091 | -34.0971 | -9.0872
5.0971 | -34.0974 | -9.0972
5.1032 | -34.0977 | -9.1078
5.1093 | -34.0979 | -9.1181
5.1154 | -34.0982 | -9.1281
5.1215 | -34.0984 | -9.1382
5.1276 | -34.0987 | -9.1483
5.1337 | -34.099 | -9.1583
5.1398 | -34.0992 | -9.1683
5.1459 | -34.0995 | -9.1788
5.152 | -34.0997 | -9.1892
5.1581 | -34.1 | -9.1997
5.1642 | -34.1003 | -9.21
5.1703 | -34.1005 | -9.2203
5.1764 | -34.1009 | -9.2305
5.1825 | -34.1011 | -9.2405
5.1886 | -34.1014 | -9.2504
5.1947 | -34.1016 | -9.2607
5.2008 | -34.1019 | -9.2711
5.2069 | -34.1022 | -9.2817
5.213 | -34.1025 | -9.2917
5.2191 | -34.1028 | -9.3021
5.2252 | -34.103 | -9.3124
5.2314 | -34.1033 | -9.3229
5.2375 | -34.1035 | -9.3329
5.2436 | -34.1038 | -9.3433
5.2497 | -34.1041 | -9.3534
5.2558 | -34.1044 | -9.364
5.2619 | -34.1047 | -9.3746
5.268 | -34.105 | -9.385
5.2741 | -34.1053 | -9.3954
5.2802 | -34.1055 | -9.4061
5.2863 | -34.1059 | -9.4163
5.2924 | -34.1061 | -9.4268
5.2985 | -34.1065 | -9.4372
5.3046 | -34.1067 | -9.4476
5.3107 | -34.107 | -9.4581
5.3168 | -34.1073 | -9.4686
5.3229 | -34.1076 | -9.4789
5.329 | -34.1079 | -9.4893
5.3351 | -34.1082 | -9.4993
5.3412 | -34.1084 | -9.5097
5.3473 | -34.1087 | -9.5201
5.3534 | -34.109 | -9.5306
5.3595 | -34.1093 | -9.5411
5.3656 | -34.1096 | -9.5514
5.3717 | -34.1099 | -9.5621
5.3779 | -34.1102 | -9.5729
5.384 | -34.1105 | -9.5832
5.3901 | -34.1108 | -9.5938
5.3962 | -34.1111 | -9.6041
5.4023 | -34.1114 | -9.6147
5.4084 | -34.1117 | -9.625
5.4145 | -34.112 | -9.6352
5.4206 | -34.1123 | -9.6456
5.4267 | -34.1126 | -9.6563
5.4328 | -34.1129 | -9.6668
5.4389 | -34.1132 | -9.677
5.445 | -34.1134 | -9.6877
5.4511 | -34.1138 | -9.698
5.4572 | -34.1141 | -9.7085
5.4633 | -34.1144 | -9.7189
5.4694 | -34.1146 | -9.7294
5.4755 | -34.1149 | -9.7395
5.4816 | -34.1152 | -9.75
5.4877 | -34.1155 | -9.7603
5.4938 | -34.1158 | -9.7705
5.4999 | -34.1161 | -9.7808
5.506 | -34.1165 | -9.7912
5.5121 | -34.1167 | -9.8015
5.5183 | -34.117 | -9.8118
5.5244 | -34.1173 | -9.8225
5.5305 | -34.1176 | -9.8328
5.5366 | -34.1179 | -9.843
5.5427 | -34.1182 | -9.8536
5.5488 | -34.1185 | -9.8641
5.5549 | -34.1187 | -9.8746
5.561 | -34.1191 | -9.8852
5.5671 | -34.1194 | -9.8955
5.5732 | -34.1197 | -9.9062
5.5793 | -34.12 | -9.917
5.5854 | -34.1203 | -9.9274
5.5915 | -34.1206 | -9.9382
5.5976 | -34.1209 | -9.9487
5.6037 | -34.1212 | -9.9593
5.6098 | -34.1214 | -9.9697
5.6159 | -34.1218 | -9.9799
5.622 | -34.122 | -9.9905
5.6281 | -34.1224 | -10.0011
5.6342 | -34.1227 | -10.0118
5.6403 | -34.123 | -10.0223
5.6464 | -34.1233 | -10.0329
5.6525 | -34.1235 | -10.0435
5.6586 | -34.1239 | -10.0537
5.6648 | -34.1242 | -10.0642
5.6709 | -34.1245 | -10.0749
5.677 | -34.1248 | -10.0854
5.6831 | -34.1251 | -10.0959
5.6892 | -34.1254 | -10.1066
5.6953 | -34.1257 | -10.1171
5.7014 | -34.126 | -10.1277
5.7075 | -34.1263 | -10.1382
5.7136 | -34.1266 | -10.1487
5.7197 | -34.1269 | -10.1592
5.7258 | -34.1271 | -10.1695
5.7319 | -34.1275 | -10.1801
5.738 | -34.1278 | -10.1905
5.7441 | -34.1281 | -10.201
5.7502 | -34.1284 | -10.2112
5.7563 | -34.1287 | -10.2217
5.7624 | -34.1291 | -10.2324
5.7685 | -34.1294 | -10.243
5.7746 | -34.1296 | -10.2537
5.7807 | -34.1299 | -10.2643
5.7868 | -34.1302 | -10.275
5.7929 | -34.1305 | -10.2854
5.799 | -34.1308 | -10.2958
5.8052 | -34.131 | -10.3061
5.8113 | -34.1313 | -10.3166
5.8174 | -34.1316 | -10.327
5.8235 | -34.1319 | -10.3374
5.8296 | -34.1322 | -10.3481
5.8357 | -34.1325 | -10.3585
5.8418 | -34.1328 | -10.3692
5.8479 | -34.1331 | -10.3797
5.854 | -34.1334 | -10.3904
5.8601 | -34.1337 | -10.4007
5.8662 | -34.134 | -10.4114
5.8723 | -34.1342 | -10.4217
5.8784 | -34.1345 | -10.4321
5.8845 | -34.1348 | -10.4425
5.8906 | -34.1351 | -10.4531
5.8967 | -34.1354 | -10.4634
5.9028 | -34.1357 | -10.4737
5.9089 | -34.1361 | -10.484
5.915 | -34.1363 | -10.4943
5.9211 | -34.1366 | -10.5047
5.9272 | -34.1369 | -10.515
5.9333 | -34.1372 | -10.5253
5.9394 | -34.1375 | -10.5358
5.9455 | -34.1378 | -10.546
5.9517 | -34.1382 | -10.5563
5.9578 | -34.1385 | -10.5669
5.9639 | -34.1388 | -10.577
5.97 | -34.1391 | -10.5875
5.9761 | -34.1395 | -10.5979
5.9822 | -34.1397 | -10.6083
5.9883 | -34.1401 | -10.6188
5.9944 | -34.1404 | -10.6293
6.0005 | -34.1407 | -10.6399
6.0066 | -34.141 | -10.6501
6.0127 | -34.1413 | -10.6603
6.0188 | -34.1416 | -10.671
6.0249 | -34.1419 | -10.6815
6.031 | -34.1422 | -10.6918
6.0371 | -34.1425 | -10.7022
6.0432 | -34.1428 | -10.7125
6.0493 | -34.1431 | -10.7227
6.0554 | -34.1434 | -10.733
6.0615 | -34.1437 | -10.7434
6.0676 | -34.1441 | -10.754
6.0737 | -34.1444 | -10.7645
6.0798 | -34.1446 | -10.7749
6.0859 | -34.145 | -10.7852
6.0921 | -34.1453 | -10.7953
6.0982 | -34.1456 | -10.8057
6.1043 | -34.1459 | -10.816
6.1104 | -34.1463 | -10.8266
6.1165 | -34.1466 | -10.8365
6.1226 | -34.1469 | -10.8468
6.1287 | -34.1472 | -10.8571
6.1348 | -34.1475 | -10.8674
6.1409 | -34.1478 | -10.8777
6.147 | -34.1481 | -10.888
6.1531 | -34.1484 | -10.8983
6.1592 | -34.1487 | -10.9085
6.1653 | -34.149 | -10.9187
6.1714 | -34.1493 | -10.929
6.1775 | -34.1496 | -10.9395
6.1836 | -34.1499 | -10.9498
6.1897 | -34.1503 | -10.9601
6.1958 | -34.1505 | -10.9707
6.2019 | -34.1509 | -10.9811
6.208 | -34.1512 | -10.9915
6.2141 | -34.1515 | -11.0017
6.2202 | -34.1519 | -11.0121
6.2263 | -34.1522 | -11.0223
6.2325 | -34.1524 | -11.0323
6.2386 | -34.1528 | -11.0424
6.2447 | -34.1531 | -11.0529
6.2508 | -34.1534 | -11.0632
6.2569 | -34.1538 | -11.0734
6.263 | -34.1541 | -11.0837
6.2691 | -34.1544 | -11.0939
6.2752 | -34.1546 | -11.1042
6.2813 | -34.1549 | -11.1145
6.2874 | -34.1552 | -11.1247
6.2935 | -34.1555 | -11.1348
6.2996 | -34.1558 | -11.1451
6.3057 | -34.1561 | -11.1554
6.3118 | -34.1564 | -11.1658
6.3179 | -34.1567 | -11.176
6.324 | -34.157 | -11.1862
6.3301 | -34.1573 | -11.1964
6.3362 | -34.1577 | -11.2066
6.3423 | -34.158 | -11.217
6.3484 | -34.1583 | -11.2273
6.3545 | -34.1586 | -11.2373
6.3606 | -34.1589 | -11.2473
6.3667 | -34.1591 | -11.2573
6.3728 | -34.1595 | -11.2676
6.379 | -34.1598 | -11.2778
6.3851 | -34.1601 | -11.2878
6.3912 | -34.1604 | -11.298
6.3973 | -34.1607 | -11.3082
6.4034 | -34.161 | -11.3182
6.4095 | -34.1613 | -11.3284
6.4156 | -34.1616 | -11.3386
6.4217 | -34.1619 | -11.3488
6.4278 | -34.1622 | -11.3591
6.4339 | -34.1625 | -11.3692
6.44 | -34.1629 | -11.3792
6.4461 | -34.1631 | -11.3897
6.4522 | -34.1635 | -11.3996
6.4583 | -34.1637 | -11.4099
6.4644 | -34.164 | -11.4202
6.4705 | -34.1643 | -11.43
6.4766 | -34.1646 | -11.4402
6.4827 | -34.1649 | -11.4505
6.4888 | -34.1652 | -11.4605
6.4949 | -34.1655 | -11.4704
6.501 | -34.1658 | -11.4806
6.5071 | -34.1661 | -11.4905
6.5132 | -34.1665 | -11.5007
6.5194 | -34.1667 | -11.5109
6.5255 | -34.1671 | -11.5208
6.5316 | -34.1673 | -11.5311
6.5377 | -34.1676 | -11.5412
6.5438 | -34.1679 | -11.5511
6.5499 | -34.1682 | -11.5611
6.556 | -34.1685 | -11.5712
6.5621 | -34.1688 | -11.5816
6.5682 | -34.1691 | -11.5917
6.5743 | -34.1695 | -11.6017
6.5804 | -34.1698 | -11.6119
6.5865 | -34.1701 | -11.6221
6.5926 | -34.1704 | -11.6321
6.5987 | -34.1707 | -11.6419
6.6048 | -34.171 | -11.652
6.6109 | -34.1713 | -11.6621
6.617 | -34.1716 | -11.6723
6.6231 | -34.1719 | -11.6824
6.6292 | -34.1722 | -11.6924
6.6353 | -34.1725 | -11.7028
6.6414 | -34.1728 | -11.7128
6.6475 | -34.1732 | -11.7229
6.6536 | -34.1734 | -11.7328
6.6597 | -34.1737 | -11.7426
6.6659 | -34.174 | -11.7527
6.672 | -34.1744 | -11.7627
6.6781 | -34.1747 | -11.7728
6.6842 | -34.175 | -11.7827
6.6903 | -34.1753 | -11.7927
6.6964 | -34.1757 | -11.8028
6.7025 | -34.1759 | -11.8126
6.7086 | -34.1762 | -11.8227
6.7147 | -34.1765 | -11.8328
6.7208 | -34.1768 | -11.8429
6.7269 | -34.1771 | -11.8529
6.733 | -34.1774 | -11.8628
6.7391 | -34.1777 | -11.8728
6.7452 | -34.1781 | -11.8826
6.7513 | -34.1784 | -11.8923
6.7574 | -34.1787 | -11.9024
6.7635 | -34.179 | -11.9123
6.7696 | -34.1793 | -11.9221
6.7757 | -34.1796 | -11.9322
6.7818 | -34.18 | -11.9421
6.7879 | -34.1803 | -11.952
6.794 | -34.1806 | -11.9618
6.8001 | -34.1809 | -11.9717
6.8063 | -34.1812 | -11.9816
6.8124 | -34.1815 | -11.9918
6.8185 | -34.1818 | -12.0019
6.8246 | -34.1821 | -12.0117
6.8307 | -34.1824 | -12.0216
6.8368 | -34.1827 | -12.0318
6.8429 | -34.183 | -12.0418
6.849 | -34.1833 | -12.0518
6.8551 | -34.1837 | -12.0616
6.8612 | -34.184 | -12.0716
6.8673 | -34.1843 | -12.0814
6.8734 | -34.1846 | -12.0915
6.8795 | -34.1849 | -12.1015
6.8856 | -34.1852 | -12.1115
6.8917 | -34.1855 | -12.1214
6.8978 | -34.1858 | -12.131
6.9039 | -34.1861 | -12.1407
6.91 | -34.1865 | -12.1507
6.9161 | -34.1868 | -12.1604
6.9222 | -34.1871 | -12.1703
6.9283 | -34.1874 | -12.1799
6.9344 | -34.1877 | -12.1899
6.9405 | -34.1881 | -12.1998
6.9466 | -34.1884 | -12.2099
6.9528 | -34.1887 | -12.22
6.9589 | -34.1891 | -12.23
6.965 | -34.1894 | -12.24
6.9711 | -34.1898 | -12.25
6.9772 | -34.1901 | -12.2599
6.9833 | -34.1904 | -12.2698
6.9894 | -34.1907 | -12.2799
6.9955 | -34.191 | -12.2896
7.0016 | -34.1914 | -12.2995
7.0077 | -34.1917 | -12.3094
7.0138 | -34.192 | -12.3192
7.0199 | -34.1923 | -12.329
7.026 | -34.1926 | -12.3388
7.0321 | -34.193 | -12.3485
7.0382 | -34.1933 | -12.358
7.0443 | -34.1937 | -12.3681
7.0504 | -34.194 | -12.3781
7.0565 | -34.1943 | -12.388
7.0626 | -34.1946 | -12.3976
7.0687 | -34.1949 | -12.407
7.0748 | -34.1953 | -12.4167
7.0809 | -34.1956 | -12.4265
7.087 | -34.1959 | -12.4364
7.0932 | -34.1962 | -12.4463
7.0993 | -34.1966 | -12.4561
7.1054 | -34.1969 | -12.4658
7.1115 | -34.1973 | -12.4755
7.1176 | -34.1976 | -12.4852
7.1237 | -34.1979 | -12.495
7.1298 | -34.1982 | -12.5048
7.1359 | -34.1985 | -12.5146
7.142 | -34.1989 | -12.5242
7.1481 | -34.1992 | -12.5339
7.1542 | -34.1995 | -12.5438
7.1603 | -34.1998 | -12.5535
7.1664 | -34.2002 | -12.5633
7.1725 | -34.2005 | -12.5732
7.1786 | -34.2009 | -12.583
7.1847 | -34.2012 | -12.5928
7.1908 | -34.2015 | -12.6028
7.1969 | -34.2018 | -12.6128
7.203 | -34.2021 | -12.6224
7.2091 | -34.2025 | -12.6321
7.2152 | -34.2028 | -12.6419
7.2213 | -34.2031 | -12.6516
7.2274 | -34.2035 | -12.6613
7.2335 | -34.2038 | -12.671
7.2397 | -34.2041 | -12.6808
7.2458 | -34.2045 | -12.6904
7.2519 | -34.2048 | -12.7003
7.258 | -34.2052 | -12.7102
7.2641 | -34.2055 | -12.72
7.2702 | -34.2058 | -12.73
7.2763 | -34.2062 | -12.7398
7.2824 | -34.2065 | -12.7496
7.2885 | -34.2068 | -12.7595
7.2946 | -34.2072 | -12.7692
7.3007 | -34.2075 | -12.7791
7.3068 | -34.2078 | -12.7891
7.3129 | -34.2081 | -12.7991
7.319 | -34.2085 | -12.8089
7.3251 | -34.2088 | -12.8185
7.3312 | -34.2091 | -12.8283
7.3373 | -34.2095 | -12.8379
7.3434 | -34.2098 | -12.8476
7.3495 | -34.2102 | -12.8573
7.3556 | -34.2105 | -12.8672
7.3617 | -34.2109 | -12.8771
7.3678 | -34.2112 | -12.8868
7.3739 | -34.2115 | -12.8969
7.3801 | -34.2118 | -12.9069
7.3862 | -34.2122 | -12.9166
7.3923 | -34.2126 | -12.9266
7.3984 | -34.2129 | -12.9361
7.4045 | -34.2133 | -12.9459
7.4106 | -34.2136 | -12.9556
7.4167 | -34.2139 | -12.9654
7.4228 | -34.2143 | -12.9753
7.4289 | -34.2146 | -12.985
7.435 | -34.215 | -12.9946
7.4411 | -34.2153 | -13.0045
7.4472 | -34.2156 | -13.0144
7.4533 | -34.216 | -13.024
7.4594 | -34.2163 | -13.034
7.4655 | -34.2167 | -13.0435
7.4716 | -34.217 | -13.0535
7.4777 | -34.2173 | -13.0632
7.4838 | -34.2177 | -13.0729
7.4899 | -34.218 | -13.0826
7.496 | -34.2184 | -13.0922
7.5021 | -34.2187 | -13.102
7.5082 | -34.219 | -13.1117
7.5143 | -34.2194 | -13.1216
7.5204 | -34.2197 | -13.1315
7.5266 | -34.2201 | -13.1412
7.5327 | -34.2204 | -13.1508
7.5388 | -34.2207 | -13.1609
7.5449 | -34.2211 | -13.1709
7.551 | -34.2214 | -13.1806
7.5571 | -34.2218 | -13.1905
7.5632 | -34.2222 | -13.2004
7.5693 | -34.2225 | -13.2101
7.5754 | -34.2229 | -13.2197
7.5815 | -34.2232 | -13.2295
7.5876 | -34.2235 | -13.2395
7.5937 | -34.2239 | -13.2495
7.5998 | -34.2243 | -13.2593
7.6059 | -34.2247 | -13.2692
7.612 | -34.225 | -13.279
7.6181 | -34.2253 | -13.2888
7.6242 | -34.2257 | -13.2988
7.6303 | -34.226 | -13.3088
7.6364 | -34.2263 | -13.3187
7.6425 | -34.2267 | -13.3286
7.6486 | -34.227 | -13.3384
7.6547 | -34.2274 | -13.3481
7.6608 | -34.2277 | -13.3579
7.667 | -34.2281 | -13.3678
7.6731 | -34.2284 | -13.3777
7.6792 | -34.2288 | -13.3877
7.6853 | -34.2291 | -13.3976
7.6914 | -34.2295 | -13.4075
7.6975 | -34.2298 | -13.4174
7.7036 | -34.2302 | -13.4273
7.7097 | -34.2305 | -13.4373
7.7158 | -34.2309 | -13.4472
7.7219 | -34.2313 | -13.4569
7.728 | -34.2316 | -13.4669
7.7341 | -34.2319 | -13.4768
7.7402 | -34.2323 | -13.4867
7.7463 | -34.2327 | -13.4966
7.7524 | -34.233 | -13.5067
7.7585 | -34.2333 | -13.5166
7.7646 | -34.2337 | -13.5264
7.7707 | -34.234 | -13.5365
7.7768 | -34.2343 | -13.5466
7.7829 | -34.2347 | -13.5567
7.789 | -34.235 | -13.5668
7.7951 | -34.2354 | -13.5767
7.8012 | -34.2357 | -13.5866
7.8073 | -34.2361 | -13.5967
7.8135 | -34.2365 | -13.6065
7.8196 | -34.2369 | -13.6167
7.8257 | -34.2372 | -13.6268
7.8318 | -34.2375 | -13.6368
7.8379 | -34.2379 | -13.6466
7.844 | -34.2382 | -13.6564
7.8501 | -34.2386 | -13.6665
7.8562 | -34.2389 | -13.6767
7.8623 | -34.2393 | -13.6865
7.8684 | -34.2397 | -13.6965
7.8745 | -34.24 | -13.7064
7.8806 | -34.2403 | -13.7165
7.8867 | -34.2407 | -13.7262
7.8928 | -34.2411 | -13.7359
7.8989 | -34.2414 | -13.7458
7.905 | -34.2418 | -13.7558
7.9111 | -34.2421 | -13.7658
7.9172 | -34.2425 | -13.7757
7.9233 | -34.2428 | -13.7857
7.9294 | -34.2432 | -13.7956
7.9355 | -34.2435 | -13.8053
7.9416 | -34.2438 | -13.8153
7.9477 | -34.2442 | -13.8251
7.9539 | -34.2445 | -13.835
7.96 | -34.2449 | -13.845
7.9661 | -34.2453 | -13.855
7.9722 | -34.2457 | -13.8648
7.9783 | -34.246 | -13.875
7.9844 | -34.2464 | -13.8849
7.9905 | -34.2467 | -13.895
7.9966 | -34.2471 | -13.905
8.0027 | -34.2474 | -13.915
8.0088 | -34.2477 | -13.9248
8.0149 | -34.2481 | -13.9348
8.021 | -34.2484 | -13.9449
8.0271 | -34.2488 | -13.9551
8.0332 | -34.2491 | -13.965
8.0393 | -34.2495 | -13.9751
8.0454 | -34.2499 | -13.9853
8.0515 | -34.2503 | -13.9952
8.0576 | -34.2506 | -14.0054
8.0637 | -34.2509 | -14.0153
8.0698 | -34.2513 | -14.0251
8.0759 | -34.2516 | -14.035
8.082 | -34.252 | -14.0448
8.0881 | -34.2523 | -14.0548
8.0942 | -34.2527 | -14.0648
8.1004 | -34.2531 | -14.0749
8.1065 | -34.2535 | -14.0847
8.1126 | -34.2538 | -14.0947
8.1187 | -34.2541 | -14.1047
8.1248 | -34.2545 | -14.1145
8.1309 | -34.2548 | -14.1243
8.137 | -34.2552 | -14.1341
8.1431 | -34.2555 | -14.1442
8.1492 | -34.2559 | -14.1541
8.1553 | -34.2562 | -14.1641
8.1614 | -34.2566 | -14.1742
8.1675 | -34.257 | -14.1842
8.1736 | -34.2573 | -14.1944
8.1797 | -34.2576 | -14.2044
8.1858 | -34.258 | -14.2144
8.1919 | -34.2584 | -14.2243
8.198 | -34.2587 | -14.2341
8.2041 | -34.2591 | -14.2441
8.2102 | -34.2595 | -14.254
8.2163 | -34.2598 | -14.2638
8.2224 | -34.2602 | -14.2738
8.2285 | -34.2605 | -14.2835
8.2346 | -34.2609 | -14.2934
8.2408 | -34.2613 | -14.3031
8.2469 | -34.2616 | -14.3132
8.253 | -34.262 | -14.323
8.2591 | -34.2624 | -14.333
8.2652 | -34.2628 | -14.343
8.2713 | -34.2631 | -14.353
8.2774 | -34.2635 | -14.3631
8.2835 | -34.2638 | -14.3733
8.2896 | -34.2642 | -14.383
8.2957 | -34.2646 | -14.393
8.3018 | -34.2649 | -14.403
8.3079 | -34.2653 | -14.4127
8.314 | -34.2657 | -14.4227
8.3201 | -34.266 | -14.4325
8.3262 | -34.2664 | -14.4426
8.3323 | -34.2667 | -14.4524
8.3384 | -34.2671 | -14.4625
8.3445 | -34.2674 | -14.4724
8.3506 | -34.2678 | -14.4823
8.3567 | -34.2682 | -14.4924
8.3628 | -34.2686 | -14.5023
8.3689 | -34.2689 | -14.5121
8.375 | -34.2693 | -14.5222
8.3812 | -34.2696 | -14.5319
8.3873 | -34.27 | -14.5418
8.3934 | -34.2703 | -14.5519
8.3995 | -34.2707 | -14.5617
8.4056 | -34.2711 | -14.5716
8.4117 | -34.2714 | -14.5816
8.4178 | -34.2718 | -14.5916
8.4239 | -34.2721 | -14.6014
8.43 | -34.2725 | -14.6113
8.4361 | -34.2728 | -14.6214
8.4422 | -34.2732 | -14.6312
8.4483 | -34.2736 | -14.641
8.4544 | -34.2739 | -14.6509
8.4605 | -34.2743 | -14.6607
8.4666 | -34.2747 | -14.6708
8.4727 | -34.275 | -14.6808
8.4788 | -34.2754 | -14.6907
8.4849 | -34.2757 | -14.7005
8.491 | -34.2761 | -14.7105
8.4971 | -34.2765 | -14.7202
8.5032 | -34.2768 | -14.73
8.5093 | -34.2772 | -14.7398
8.5154 | -34.2776 | -14.7498
8.5215 | -34.2779 | -14.7596
8.5277 | -34.2783 | -14.7694
8.5338 | -34.2786 | -14.7793
8.5399 | -34.279 | -14.7891
8.546 | -34.2794 | -14.7988
8.5521 | -34.2797 | -14.8087
8.5582 | -34.2801 | -14.8185
8.5643 | -34.2805 | -14.8281
8.5704 | -34.2808 | -14.8377
8.5765 | -34.2812 | -14.8477
8.5826 | -34.2815 | -14.8573
8.5887 | -34.2819 | -14.8672
8.5948 | -34.2823 | -14.877
8.6009 | -34.2827 | -14.8867
8.607 | -34.283 | -14.8967
8.6131 | -34.2834 | -14.9066
8.6192 | -34.2838 | -14.9162
8.6253 | -34.2842 | -14.926
8.6314 | -34.2845 | -14.9357
8.6375 | -34.2849 | -14.9456
8.6436 | -34.2852 | -14.9555
8.6497 | -34.2856 | -14.9652
8.6558 | -34.2859 | -14.975
8.6619 | -34.2863 | -14.9848
8.6681 | -34.2866 | -14.9948
8.6742 | -34.287 | -15.0045
8.6803 | -34.2874 | -15.0144
8.6864 | -34.2877 | -15.0239
8.6925 | -34.2881 | -15.0335
8.6986 | -34.2884 | -15.0434
8.7047 | -34.2888 | -15.0534
8.7108 | -34.2891 | -15.0632
8.7169 | -34.2895 | -15.0729
8.723 | -34.2899 | -15.0825
8.7291 | -34.2902 | -15.0921
8.7352 | -34.2906 | -15.1021
8.7413 | -34.291 | -15.1119
8.7474 | -34.2913 | -15.1215
8.7535 | -34.2917 | -15.1312
8.7596 | -34.2921 | -15.1411
8.7657 | -34.2925 | -15.151
8.7718 | -34.2928 | -15.1608
8.7779 | -34.2932 | -15.1707
8.784 | -34.2935 | -15.1806
8.7901 | -34.2939 | -15.1903
8.7962 | -34.2943 | -15.1998
8.8023 | -34.2947 | -15.2096
8.8084 | -34.295 | -15.2193
8.8146 | -34.2954 | -15.229
8.8207 | -34.2957 | -15.2387
8.8268 | -34.2962 | -15.2481
8.8329 | -34.2965 | -15.2579
8.839 | -34.2969 | -15.2676
8.8451 | -34.2973 | -15.2773
8.8512 | -34.2976 | -15.2869
8.8573 | -34.298 | -15.2966
8.8634 | -34.2984 | -15.3061
8.8695 | -34.2987 | -15.316
8.8756 | -34.2991 | -15.3257
8.8817 | -34.2994 | -15.3353
8.8878 | -34.2998 | -15.3453
8.8939 | -34.3002 | -15.3551
8.9 | -34.3005 | -15.3648
8.9061 | -34.3009 | -15.3744
8.9122 | -34.3013 | -15.3842
8.9183 | -34.3017 | -15.3939
8.9244 | -34.3021 | -15.4037
8.9305 | -34.3025 | -15.4131
8.9366 | -34.3028 | -15.4228
8.9427 | -34.3032 | -15.4326
8.9488 | -34.3036 | -15.4422
8.955 | -34.3039 | -15.4518
8.9611 | -34.3043 | -15.4617
8.9672 | -34.3047 | -15.4712
8.9733 | -34.3051 | -15.4807
8.9794 | -34.3055 | -15.4903
8.9855 | -34.3059 | -15.5001
8.9916 | -34.3062 | -15.5099
8.9977 | -34.3066 | -15.5195
9.0038 | -34.307 | -15.529
9.0099 | -34.3074 | -15.5387
9.016 | -34.3078 | -15.5484
9.0221 | -34.3082 | -15.558
9.0282 | -34.3085 | -15.5674
9.0343 | -34.3089 | -15.5772
9.0404 | -34.3093 | -15.5868
9.0465 | -34.3097 | -15.5967
9.0526 | -34.3101 | -15.6062
9.0587 | -34.3105 | -15.616
9.0648 | -34.3109 | -15.6255
9.0709 | -34.3113 | -15.6352
9.077 | -34.3117 | -15.6449
9.0831 | -34.3121 | -15.6547
9.0892 | -34.3124 | -15.6645
9.0953 | -34.3128 | -15.6742
9.1015 | -34.3132 | -15.6838
9.1076 | -34.3136 | -15.6933
9.1137 | -34.314 | -15.7029
9.1198 | -34.3144 | -15.7125
9.1259 | -34.3148 | -15.7222
9.132 | -34.3152 | -15.7319
9.1381 | -34.3156 | -15.7414
9.1442 | -34.316 | -15.7507
9.1503 | -34.3164 | -15.7604
9.1564 | -34.3168 | -15.77
9.1625 | -34.3171 | -15.7797
9.1686 | -34.3176 | -15.7893
9.1747 | -34.3179 | -15.7989
9.1808 | -34.3184 | -15.8085
9.1869 | -34.3188 | -15.8181
9.193 | -34.3192 | -15.8277
9.1991 | -34.3195 | -15.8374
9.2052 | -34.3199 | -15.847
9.2113 | -34.3203 | -15.8567
9.2174 | -34.3208 | -15.8663
9.2235 | -34.3211 | -15.8759
9.2296 | -34.3216 | -15.8855
9.2357 | -34.3219 | -15.8952
9.2419 | -34.3223 | -15.9048
9.248 | -34.3228 | -15.9143
9.2541 | -34.3232 | -15.924
9.2602 | -34.3236 | -15.9337
9.2663 | -34.324 | -15.9432
9.2724 | -34.3244 | -15.9527
9.2785 | -34.3248 | -15.9623
9.2846 | -34.3252 | -15.9718
9.2907 | -34.3256 | -15.9814
9.2968 | -34.326 | -15.9911
9.3029 | -34.3264 | -16.0008
9.309 | -34.3268 | -16.0104
9.3151 | -34.3271 | -16.0199
9.3212 | -34.3276 | -16.0295
9.3273 | -34.328 | -16.039
9.3334 | -34.3284 | -16.0485
9.3395 | -34.3287 | -16.0579
9.3456 | -34.3291 | -16.0674
9.3517 | -34.3295 | -16.077
9.3578 | -34.3299 | -16.0865
9.3639 | -34.3304 | -16.0959
9.37 | -34.3308 | -16.1056
9.3761 | -34.3312 | -16.1153
9.3822 | -34.3316 | -16.1249
9.3884 | -34.332 | -16.1342
9.3945 | -34.3324 | -16.1438
9.4006 | -34.3328 | -16.1534
9.4067 | -34.3332 | -16.1631
9.4128 | -34.3336 | -16.1728
9.4189 | -34.334 | -16.1825
9.425 | -34.3345 | -16.192
9.4311 | -34.3349 | -16.2015
9.4372 | -34.3353 | -16.2111
9.4433 | -34.3357 | -16.2207
9.4494 | -34.3361 | -16.23
9.4555 | -34.3365 | -16.2397
9.4616 | -34.3369 | -16.2492
9.4677 | -34.3373 | -16.2588
9.4738 | -34.3377 | -16.2683
9.4799 | -34.3381 | -16.2778
9.486 | -34.3385 | -16.2874
9.4921 | -34.3389 | -16.297
9.4982 | -34.3394 | -16.3064
9.5043 | -34.3398 | -16.3159
9.5104 | -34.3402 | -16.3253
9.5165 | -34.3406 | -16.3349
9.5226 | -34.341 | -16.3445
9.5288 | -34.3414 | -16.3538
9.5349 | -34.3418 | -16.3634
9.541 | -34.3422 | -16.3728
9.5471 | -34.3426 | -16.3823
9.5532 | -34.343 | -16.3918
9.5593 | -34.3435 | -16.4014
9.5654 | -34.3439 | -16.4108
9.5715 | -34.3443 | -16.4203
9.5776 | -34.3447 | -16.43
9.5837 | -34.3451 | -16.4395
9.5898 | -34.3455 | -16.4491
9.5959 | -34.3459 | -16.4587
9.602 | -34.3463 | -16.4681
9.6081 | -34.3467 | -16.4776
9.6142 | -34.3471 | -16.487
9.6203 | -34.3475 | -16.4965
9.6264 | -34.3479 | -16.506
9.6325 | -34.3483 | -16.5156
9.6386 | -34.3487 | -16.525
9.6447 | -34.3491 | -16.5346
9.6508 | -34.3495 | -16.544
9.6569 | -34.3499 | -16.5535
9.663 | -34.3503 | -16.5631
9.6691 | -34.3507 | -16.5726
9.6753 | -34.3512 | -16.5821
9.6814 | -34.3516 | -16.5915
9.6875 | -34.352 | -16.6008
9.6936 | -34.3524 | -16.6102
9.6997 | -34.3528 | -16.6199
9.7058 | -34.3532 | -16.6292
9.7119 | -34.3536 | -16.6388
9.718 | -34.3541 | -16.6484
9.7241 | -34.3545 | -16.6578
9.7302 | -34.3549 | -16.6674
9.7363 | -34.3553 | -16.6768
9.7424 | -34.3557 | -16.6861
9.7485 | -34.3561 | -16.6957
9.7546 | -34.3565 | -16.7051
9.7607 | -34.3569 | -16.7148
9.7668 | -34.3573 | -16.7243
9.7729 | -34.3578 | -16.7338
9.779 | -34.3582 | -16.743
9.7851 | -34.3586 | -16.7527
9.7912 | -34.3591 | -16.7621
9.7973 | -34.3595 | -16.7717
9.8034 | -34.3599 | -16.781
9.8095 | -34.3603 | -16.7908
9.8157 | -34.3608 | -16.8003
9.8218 | -34.3612 | -16.81
9.8279 | -34.3616 | -16.8194
9.834 | -34.362 | -16.829
9.8401 | -34.3624 | -16.8384
9.8462 | -34.3628 | -16.8478
9.8523 | -34.3632 | -16.8573
9.8584 | -34.3636 | -16.8667
9.8645 | -34.364 | -16.8763
9.8706 | -34.3644 | -16.8858
9.8767 | -34.3648 | -16.8952
9.8828 | -34.3652 | -16.9047
9.8889 | -34.3656 | -16.914
9.895 | -34.366 | -16.9234
9.9011 | -34.3665 | -16.9328
9.9072 | -34.3669 | -16.9423
9.9133 | -34.3673 | -16.9518
9.9194 | -34.3677 | -16.9612
9.9255 | -34.3681 | -16.9706
9.9316 | -34.3685 | -16.98
9.9377 | -34.3689 | -16.9895
9.9438 | -34.3693 | -16.999
9.9499 | -34.3698 | -17.0087
9.956 | -34.3702 | -17.0181
9.9622 | -34.3706 | -17.0276
9.9683 | -34.371 | -17.037
9.9744 | -34.3714 | -17.0466
9.9805 | -34.3718 | -17.056
9.9866 | -34.3722 | -17.0654
9.9927 | -34.3726 | -17.0748
9.9988 | -34.373 | -17.0843
10.0049 | -34.3734 | -17.0937
10.011 | -34.3739 | -17.1032
10.0171 | -34.3743 | -17.1125
10.0232 | -34.3747 | -17.122
10.0293 | -34.3751 | -17.1316
10.0354 | -34.3756 | -17.141
10.0415 | -34.376 | -17.1504
10.0476 | -34.3764 | -17.16
10.0537 | -34.3768 | -17.1696
10.0598 | -34.3772 | -17.179
10.0659 | -34.3776 | -17.1884
10.072 | -34.378 | -17.1977
10.0781 | -34.3784 | -17.2071
10.0842 | -34.3789 | -17.2165
10.0903 | -34.3793 | -17.2259
10.0964 | -34.3797 | -17.2351
10.1026 | -34.3801 | -17.2446
10.1087 | -34.3805 | -17.254
10.1148 | -34.3809 | -17.2635
10.1209 | -34.3813 | -17.2729
10.127 | -34.3817 | -17.2821
10.1331 | -34.3821 | -17.2915
10.1392 | -34.3825 | -17.3008
10.1453 | -34.3829 | -17.3102
10.1514 | -34.3833 | -17.3197
10.1575 | -34.3837 | -17.3292
10.1636 | -34.3842 | -17.3384
10.1697 | -34.3846 | -17.3479
10.1758 | -34.385 | -17.3571
10.1819 | -34.3854 | -17.3664
10.188 | -34.3858 | -17.3759
10.1941 | -34.3862 | -17.3853
10.2002 | -34.3866 | -17.3947
10.2063 | -34.387 | -17.4042
10.2124 | -34.3874 | -17.4137
10.2185 | -34.3879 | -17.4231
10.2246 | -34.3883 | -17.4324
10.2307 | -34.3887 | -17.4418
10.2368 | -34.3891 | -17.451
10.2429 | -34.3895 | -17.4603
10.2491 | -34.3899 | -17.4698
10.2552 | -34.3903 | -17.4792
10.2613 | -34.3907 | -17.4886
10.2674 | -34.3911 | -17.498
10.2735 | -34.3916 | -17.5076
10.2796 | -34.392 | -17.5169
10.2857 | -34.3924 | -17.5264
10.2918 | -34.3928 | -17.5359
10.2979 | -34.3932 | -17.5453
10.304 | -34.3937 | -17.5545
10.3101 | -34.394 | -17.564
10.3162 | -34.3945 | -17.5735
10.3223 | -34.3949 | -17.5827
10.3284 | -34.3953 | -17.5921
10.3345 | -34.3957 | -17.6013
10.3406 | -34.3961 | -17.6107
10.3467 | -34.3965 | -17.6201
10.3528 | -34.3969 | -17.6294
10.3589 | -34.3973 | -17.6389
10.365 | -34.3977 | -17.6483
10.3711 | -34.3981 | -17.6578
10.3772 | -34.3985 | -17.6672
10.3833 | -34.3989 | -17.6766
10.3895 | -34.3994 | -17.686
10.3956 | -34.3998 | -17.6955
10.4017 | -34.4002 | -17.7047
10.4078 | -34.4006 | -17.714
10.4139 | -34.401 | -17.7235
10.42 | -34.4014 | -17.7331
10.4261 | -34.4018 | -17.7424
10.4322 | -34.4022 | -17.7517
10.4383 | -34.4026 | -17.761
10.4444 | -34.403 | -17.7702
10.4505 | -34.4034 | -17.7796
10.4566 | -34.4038 | -17.7888
10.4627 | -34.4042 | -17.798
10.4688 | -34.4046 | -17.8072
10.4749 | -34.405 | -17.8165
10.481 | -34.4055 | -17.826
10.4871 | -34.4059 | -17.8352
10.4932 | -34.4063 | -17.8445
10.4993 | -34.4067 | -17.8539
10.5054 | -34.4071 | -17.8632
10.5115 | -34.4075 | -17.8726
10.5176 | -34.408 | -17.8818
10.5237 | -34.4084 | -17.8913
10.5298 | -34.4088 | -17.9005
10.536 | -34.4092 | -17.91
10.5421 | -34.4096 | -17.9192
10.5482 | -34.41 | -17.9287
10.5543 | -34.4105 | -17.9379
10.5604 | -34.4109 | -17.9472
10.5665 | -34.4113 | -17.9564
10.5726 | -34.4117 | -17.9658
10.5787 | -34.4121 | -17.975
10.5848 | -34.4125 | -17.984
10.5909 | -34.4129 | -17.9934
10.597 | -34.4133 | -18.0028
10.6031 | -34.4138 | -18.0124
10.6092 | -34.4141 | -18.0217
10.6153 | -34.4145 | -18.0311
10.6214 | -34.415 | -18.0402
10.6275 | -34.4154 | -18.0496
10.6336 | -34.4158 | -18.0589
10.6397 | -34.4162 | -18.0683
10.6458 | -34.4166 | -18.0775
10.6519 | -34.417 | -18.087
10.658 | -34.4174 | -18.0964
10.6641 | -34.4178 | -18.1057
10.6702 | -34.4183 | -18.115
10.6764 | -34.4187 | -18.1245
10.6825 | -34.4191 | -18.1337
10.6886 | -34.4195 | -18.1427
10.6947 | -34.4199 | -18.1522
10.7008 | -34.4203 | -18.1614
10.7069 | -34.4208 | -18.1708
10.713 | -34.4212 | -18.1801
10.7191 | -34.4216 | -18.1894
10.7252 | -34.422 | -18.1984
10.7313 | -34.4224 | -18.2078
10.7374 | -34.4228 | -18.2171
10.7435 | -34.4232 | -18.2264
10.7496 | -34.4236 | -18.2355
10.7557 | -34.424 | -18.2449
10.7618 | -34.4244 | -18.2539
10.7679 | -34.4248 | -18.2631
10.774 | -34.4252 | -18.2724
10.7801 | -34.4257 | -18.2818
10.7862 | -34.4261 | -18.291
10.7923 | -34.4265 | -18.3004
10.7984 | -34.4269 | -18.3096
10.8045 | -34.4273 | -18.3186
10.8106 | -34.4277 | -18.328
10.8168 | -34.4281 | -18.3373
10.8229 | -34.4285 | -18.3466
10.829 | -34.4289 | -18.3558
10.8351 | -34.4294 | -18.365
10.8412 | -34.4298 | -18.3743
10.8473 | -34.4302 | -18.3834
10.8534 | -34.4307 | -18.3927
10.8595 | -34.4311 | -18.402
10.8656 | -34.4315 | -18.4111
10.8717 | -34.4319 | -18.4203
10.8778 | -34.4324 | -18.4295
10.8839 | -34.4328 | -18.4388
10.89 | -34.4332 | -18.4482
10.8961 | -34.4336 | -18.4574
10.9022 | -34.434 | -18.4666
10.9083 | -34.4344 | -18.4757
10.9144 | -34.4348 | -18.4849
10.9205 | -34.4352 | -18.4941
10.9266 | -34.4357 | -18.5033
10.9327 | -34.4361 | -18.5125
10.9388 | -34.4365 | -18.5216
10.9449 | -34.4369 | -18.5307
10.951 | -34.4374 | -18.5399
10.9571 | -34.4378 | -18.5491
10.9633 | -34.4382 | -18.5585
10.9694 | -34.4386 | -18.5679
10.9755 | -34.439 | -18.5772
10.9816 | -34.4394 | -18.5865
10.9877 | -34.4399 | -18.5957
10.9938 | -34.4403 | -18.605
10.9999 | -34.4407 | -18.6141
11.006 | -34.4412 | -18.6231
11.0121 | -34.4416 | -18.6324
11.0182 | -34.442 | -18.6415
11.0243 | -34.4425 | -18.6507
11.0304 | -34.4429 | -18.66
11.0365 | -34.4433 | -18.6691
11.0426 | -34.4437 | -18.6783
11.0487 | -34.4442 | -18.6877
11.0548 | -34.4446 | -18.6966
11.0609 | -34.445 | -18.7059
11.067 | -34.4454 | -18.7151
11.0731 | -34.4459 | -18.7242
11.0792 | -34.4463 | -18.7334
11.0853 | -34.4467 | -18.7425
11.0914 | -34.4472 | -18.7519
11.0975 | -34.4476 | -18.7611
11.1037 | -34.448 | -18.7702
11.1098 | -34.4484 | -18.7796
11.1159 | -34.4489 | -18.7885
11.122 | -34.4493 | -18.7976
11.1281 | -34.4497 | -18.8068
11.1342 | -34.4501 | -18.816
11.1403 | -34.4506 | -18.8253
11.1464 | -34.451 | -18.8347
11.1525 | -34.4514 | -18.8438
11.1586 | -34.4518 | -18.853
11.1647 | -34.4522 | -18.8621
11.1708 | -34.4526 | -18.8713
11.1769 | -34.4531 | -18.8805
11.183 | -34.4535 | -18.8895
11.1891 | -34.4539 | -18.8986
11.1952 | -34.4544 | -18.9077
11.2013 | -34.4548 | -18.9169
11.2074 | -34.4552 | -18.926
11.2135 | -34.4556 | -18.9353
11.2196 | -34.4561 | -18.9445
11.2257 | -34.4565 | -18.9538
11.2318 | -34.457 | -18.9628
11.2379 | -34.4574 | -18.9722
11.244 | -34.4579 | -18.9816
11.2502 | -34.4583 | -18.9906
11.2563 | -34.4588 | -18.9998
11.2624 | -34.4592 | -19.009
11.2685 | -34.4597 | -19.0181
11.2746 | -34.4601 | -19.0273
11.2807 | -34.4605 | -19.0365
11.2868 | -34.461 | -19.0458
11.2929 | -34.4614 | -19.0546
11.299 | -34.4619 | -19.0638
11.3051 | -34.4623 | -19.0731
11.3112 | -34.4627 | -19.0824
11.3173 | -34.4631 | -19.0914
11.3234 | -34.4636 | -19.1004
11.3295 | -34.464 | -19.1097
11.3356 | -34.4644 | -19.1189
11.3417 | -34.4649 | -19.1279
11.3478 | -34.4653 | -19.1373
11.3539 | -34.4657 | -19.1465
11.36 | -34.4661 | -19.1555
11.3661 | -34.4666 | -19.1645
11.3722 | -34.4671 | -19.1734
11.3783 | -34.4675 | -19.1826
11.3844 | -34.4679 | -19.1919
11.3906 | -34.4684 | -19.201
11.3967 | -34.4689 | -19.2101
11.4028 | -34.4693 | -19.2194
11.4089 | -34.4697 | -19.2284
11.415 | -34.4701 | -19.2375
11.4211 | -34.4706 | -19.2466
11.4272 | -34.471 | -19.2558
11.4333 | -34.4714 | -19.2646
11.4394 | -34.4719 | -19.2737
11.4455 | -34.4723 | -19.2827
11.4516 | -34.4727 | -19.2919
11.4577 | -34.4731 | -19.301
11.4638 | -34.4736 | -19.31
11.4699 | -34.4741 | -19.3192
11.476 | -34.4745 | -19.3285
11.4821 | -34.475 | -19.3374
11.4882 | -34.4754 | -19.3467
11.4943 | -34.4759 | -19.356
11.5004 | -34.4763 | -19.3648
11.5065 | -34.4768 | -19.3739
11.5126 | -34.4772 | -19.3829
11.5187 | -34.4776 | -19.392
11.5248 | -34.4781 | -19.4013
11.5309 | -34.4785 | -19.4105
11.5371 | -34.4789 | -19.4195
11.5432 | -34.4794 | -19.4286
11.5493 | -34.4798 | -19.4379
11.5554 | -34.4803 | -19.4471
11.5615 | -34.4807 | -19.4563
11.5676 | -34.4812 | -19.4655
11.5737 | -34.4816 | -19.4744
11.5798 | -34.4821 | -19.4835
11.5859 | -34.4826 | -19.4927
11.592 | -34.483 | -19.5017
11.5981 | -34.4834 | -19.5107
11.6042 | -34.4839 | -19.5198
11.6103 | -34.4843 | -19.5288
11.6164 | -34.4847 | -19.5377
11.6225 | -34.4852 | -19.5468
11.6286 | -34.4856 | -19.5558
11.6347 | -34.486 | -19.5649
11.6408 | -34.4865 | -19.574
11.6469 | -34.4869 | -19.5831
11.653 | -34.4874 | -19.592
11.6591 | -34.4878 | -19.6011
11.6652 | -34.4883 | -19.6102
11.6713 | -34.4887 | -19.6191
11.6775 | -34.4892 | -19.628
11.6836 | -34.4896 | -19.6372
11.6897 | -34.4901 | -19.6463
11.6958 | -34.4905 | -19.6555
11.7019 | -34.491 | -19.6647
11.708 | -34.4914 | -19.6737
11.7141 | -34.4919 | -19.6827
11.7202 | -34.4923 | -19.6917
11.7263 | -34.4928 | -19.7009
11.7324 | -34.4932 | -19.7101
11.7385 | -34.4937 | -19.7192
11.7446 | -34.4941 | -19.7282
11.7507 | -34.4946 | -19.7371
11.7568 | -34.495 | -19.7463
11.7629 | -34.4954 | -19.7555
11.769 | -34.4959 | -19.7644
11.7751 | -34.4963 | -19.7735
11.7812 | -34.4968 | -19.7824
11.7873 | -34.4973 | -19.7915
11.7934 | -34.4977 | -19.8007
11.7995 | -34.4982 | -19.8097
11.8056 | -34.4986 | -19.8186
11.8117 | -34.4991 | -19.8278
11.8178 | -34.4995 | -19.8367
11.824 | -34.5 | -19.8457
11.8301 | -34.5005 | -19.8546
11.8362 | -34.5009 | -19.8635
11.8423 | -34.5013 | -19.8726
11.8484 | -34.5018 | -19.8817
11.8545 | -34.5023 | -19.8904
11.8606 | -34.5027 | -19.8994
11.8667 | -34.5032 | -19.9085
11.8728 | -34.5036 | -19.9174
11.8789 | -34.504 | -19.9263
11.885 | -34.5045 | -19.9354
11.8911 | -34.505 | -19.9444
11.8972 | -34.5054 | -19.9533
11.9033 | -34.5059 | -19.9623
11.9094 | -34.5063 | -19.9712
11.9155 | -34.5068 | -19.9802
11.9216 | -34.5072 | -19.989
11.9277 | -34.5077 | -19.9979
11.9338 | -34.5081 | -20.0069
11.9399 | -34.5086 | -20.016
11.946 | -34.509 | -20.0249
11.9521 | -34.5094 | -20.0337
11.9582 | -34.5099 | -20.0426
11.9644 | -34.5104 | -20.0514
11.9705 | -34.5109 | -20.0605
11.9766 | -34.5113 | -20.0693
11.9827 | -34.5118 | -20.0783
11.9888 | -34.5123 | -20.0871
11.9949 | -34.5128 | -20.0958
12.001 | -34.5132 | -20.1049
12.0071 | -34.5137 | -20.1138
12.0132 | -34.5141 | -20.1227
12.0193 | -34.5146 | -20.1317
12.0254 | -34.515 | -20.1408
12.0315 | -34.5155 | -20.1494
12.0376 | -34.5159 | -20.1583
12.0437 | -34.5164 | -20.1673
12.0498 | -34.5168 | -20.1764
12.0559 | -34.5173 | -20.1851
12.062 | -34.5178 | -20.1939
12.0681 | -34.5182 | -20.2029
12.0742 | -34.5187 | -20.2117
12.0803 | -34.5192 | -20.2209
12.0864 | -34.5196 | -20.2299
12.0925 | -34.5201 | -20.2388
12.0986 | -34.5206 | -20.2477
12.1047 | -34.521 | -20.2567
12.1109 | -34.5215 | -20.2654
12.117 | -34.5219 | -20.2743
12.1231 | -34.5224 | -20.2831
12.1292 | -34.5229 | -20.2921
12.1353 | -34.5233 | -20.3011
12.1414 | -34.5238 | -20.31
12.1475 | -34.5242 | -20.3191
12.1536 | -34.5247 | -20.328
12.1597 | -34.5251 | -20.3369
12.1658 | -34.5256 | -20.3458
12.1719 | -34.5261 | -20.3545
12.178 | -34.5265 | -20.3635
12.1841 | -34.527 | -20.3725
12.1902 | -34.5275 | -20.3814
12.1963 | -34.5279 | -20.3903
12.2024 | -34.5284 | -20.399
12.2085 | -34.5289 | -20.4079
12.2146 | -34.5293 | -20.4168
12.2207 | -34.5298 | -20.4257
12.2268 | -34.5303 | -20.4345
12.2329 | -34.5307 | -20.4433
12.239 | -34.5312 | -20.4519
12.2451 | -34.5317 | -20.4608
12.2513 | -34.5321 | -20.4699
12.2574 | -34.5326 | -20.4788
12.2635 | -34.5331 | -20.4877
12.2696 | -34.5335 | -20.4965
12.2757 | -34.534 | -20.505
12.2818 | -34.5344 | -20.5139
12.2879 | -34.5349 | -20.5228
12.294 | -34.5354 | -20.5315
12.3001 | -34.5358 | -20.5403
12.3062 | -34.5363 | -20.5491
12.3123 | -34.5368 | -20.5576
12.3184 | -34.5372 | -20.5665
12.3245 | -34.5377 | -20.5751
12.3306 | -34.5382 | -20.584
12.3367 | -34.5386 | -20.5928
12.3428 | -34.5391 | -20.6015
12.3489 | -34.5396 | -20.6101
12.355 | -34.54 | -20.619
12.3611 | -34.5405 | -20.628
12.3672 | -34.541 | -20.6368
12.3733 | -34.5415 | -20.6457
12.3794 | -34.5419 | -20.6542
12.3855 | -34.5424 | -20.6629
12.3916 | -34.5428 | -20.6718
12.3978 | -34.5433 | -20.6807
12.4039 | -34.5438 | -20.6895
12.41 | -34.5442 | -20.6983
12.4161 | -34.5447 | -20.7069
12.4222 | -34.5451 | -20.7157
12.4283 | -34.5456 | -20.7244
12.4344 | -34.5461 | -20.7334
12.4405 | -34.5465 | -20.7423
12.4466 | -34.547 | -20.751
12.4527 | -34.5474 | -20.7599
12.4588 | -34.5479 | -20.7687
12.4649 | -34.5484 | -20.7776
12.471 | -34.5488 | -20.7866
12.4771 | -34.5493 | -20.7953
12.4832 | -34.5497 | -20.804
12.4893 | -34.5502 | -20.8127
12.4954 | -34.5507 | -20.8212
12.5015 | -34.5511 | -20.83
12.5076 | -34.5516 | -20.8388
12.5137 | -34.5521 | -20.8476
12.5198 | -34.5525 | -20.8562
12.5259 | -34.553 | -20.865
12.532 | -34.5534 | -20.8738
12.5382 | -34.5539 | -20.8827
12.5443 | -34.5543 | -20.8915
12.5504 | -34.5548 | -20.9004
12.5565 | -34.5553 | -20.909
12.5626 | -34.5558 | -20.9178
12.5687 | -34.5562 | -20.9265
12.5748 | -34.5567 | -20.9353
12.5809 | -34.5571 | -20.9439
12.587 | -34.5576 | -20.9527
12.5931 | -34.5581 | -20.9612
12.5992 | -34.5586 | -20.9699
12.6053 | -34.559 | -20.9786
12.6114 | -34.5595 | -20.9874
12.6175 | -34.56 | -20.9964
12.6236 | -34.5604 | -21.0052
12.6297 | -34.5609 | -21.0138
12.6358 | -34.5614 | -21.0223
12.6419 | -34.5618 | -21.031
12.648 | -34.5623 | -21.0398
12.6541 | -34.5628 | -21.0485
12.6602 | -34.5632 | -21.0567
12.6663 | -34.5637 | -21.0655
12.6724 | -34.5642 | -21.0741
12.6785 | -34.5646 | -21.0828
12.6847 | -34.5651 | -21.0915
12.6908 | -34.5655 | -21.1002
12.6969 | -34.566 | -21.1086
12.703 | -34.5664 | -21.1175
12.7091 | -34.5669 | -21.1263
12.7152 | -34.5674 | -21.1349
12.7213 | -34.5678 | -21.1436
12.7274 | -34.5683 | -21.1522
12.7335 | -34.5687 | -21.1607
12.7396 | -34.5692 | -21.1692
12.7457 | -34.5697 | -21.1779
12.7518 | -34.5701 | -21.1866
12.7579 | -34.5706 | -21.1953
12.764 | -34.571 | -21.2041
12.7701 | -34.5715 | -21.2127
12.7762 | -34.5719 | -21.2212
12.7823 | -34.5724 | -21.2299
12.7884 | -34.5729 | -21.2387
12.7945 | -34.5734 | -21.2476
12.8006 | -34.5738 | -21.2562
12.8067 | -34.5743 | -21.2648
12.8128 | -34.5747 | -21.2734
12.8189 | -34.5752 | -21.2821
12.8251 | -34.5756 | -21.2908
12.8312 | -34.5761 | -21.2997
12.8373 | -34.5765 | -21.3081
12.8434 | -34.577 | -21.3168
12.8495 | -34.5774 | -21.3253
12.8556 | -34.5779 | -21.3339
12.8617 | -34.5784 | -21.3425
12.8678 | -34.5788 | -21.3512
12.8739 | -34.5793 | -21.3597
12.88 | -34.5797 | -21.3682
12.8861 | -34.5802 | -21.3767
12.8922 | -34.5807 | -21.3854
12.8983 | -34.5811 | -21.3941
12.9044 | -34.5816 | -21.4025
12.9105 | -34.582 | -21.4109
12.9166 | -34.5825 | -21.4196
12.9227 | -34.5829 | -21.428
12.9288 | -34.5834 | -21.4367
12.9349 | -34.5839 | -21.4454
12.941 | -34.5843 | -21.454
12.9471 | -34.5848 | -21.4626
12.9532 | -34.5853 | -21.4712
12.9593 | -34.5857 | -21.48
12.9654 | -34.5862 | -21.4888
12.9716 | -34.5866 | -21.4976
12.9777 | -34.5871 | -21.5062
12.9838 | -34.5876 | -21.5149
12.9899 | -34.588 | -21.5236
12.996 | -34.5884 | -21.5321
13.0021 | -34.5889 | -21.5409
13.0082 | -34.5894 | -21.5494
13.0143 | -34.5898 | -21.5578
13.0204 | -34.5903 | -21.5663
13.0265 | -34.5908 | -21.5748
13.0326 | -34.5912 | -21.5834
13.0387 | -34.5917 | -21.5922
13.0448 | -34.5921 | -21.6008
13.0509 | -34.5926 | -21.6095
13.057 | -34.5931 | -21.618
13.0631 | -34.5936 | -21.6264
13.0692 | -34.594 | -21.6349
13.0753 | -34.5945 | -21.6437
13.0814 | -34.5949 | -21.6523
13.0875 | -34.5954 | -21.6606
13.0936 | -34.5958 | -21.6693
13.0997 | -34.5963 | -21.6776
13.1058 | -34.5968 | -21.6859
13.112 | -34.5972 | -21.6944
13.1181 | -34.5977 | -21.703
13.1242 | -34.5981 | -21.7116
13.1303 | -34.5986 | -21.7203
13.1364 | -34.599 | -21.7288
13.1425 | -34.5994 | -21.7371
13.1486 | -34.5998 | -21.7457
13.1547 | -34.6003 | -21.7545
13.1608 | -34.6008 | -21.7631
13.1669 | -34.6012 | -21.7715
13.173 | -34.6017 | -21.7799
13.1791 | -34.6021 | -21.7885
13.1852 | -34.6026 | -21.797
13.1913 | -34.603 | -21.805
13.1974 | -34.6035 | -21.8136
13.2035 | -34.6039 | -21.8222
13.2096 | -34.6044 | -21.8304
13.2157 | -34.6048 | -21.8388
13.2218 | -34.6053 | -21.8474
13.2279 | -34.6057 | -21.8557
13.234 | -34.6062 | -21.8642
13.2401 | -34.6066 | -21.8728
13.2462 | -34.6071 | -21.8814
13.2524 | -34.6075 | -21.8899
13.2585 | -34.6079 | -21.8983
13.2646 | -34.6084 | -21.9068
13.2707 | -34.6088 | -21.9153
13.2768 | -34.6093 | -21.924
13.2829 | -34.6098 | -21.9328
13.289 | -34.6102 | -21.9412
13.2951 | -34.6107 | -21.9492
13.3012 | -34.6111 | -21.958
13.3073 | -34.6116 | -21.9665
13.3134 | -34.612 | -21.9749
13.3195 | -34.6125 | -21.9834
13.3256 | -34.6129 | -21.9918
13.3317 | -34.6133 | -22.0002
13.3378 | -34.6138 | -22.0087
13.3439 | -34.6143 | -22.0171
13.35 | -34.6148 | -22.0259
13.3561 | -34.6152 | -22.0343
13.3622 | -34.6156 | -22.0429
13.3683 | -34.6161 | -22.0513
13.3744 | -34.6165 | -22.0599
13.3805 | -34.6169 | -22.0684
13.3866 | -34.6174 | -22.0766
13.3927 | -34.6178 | -22.0851
13.3989 | -34.6182 | -22.0935
13.405 | -34.6187 | -22.1023
13.4111 | -34.6191 | -22.1108
13.4172 | -34.6196 | -22.1191
13.4233 | -34.62 | -22.1278
13.4294 | -34.6204 | -22.1362
13.4355 | -34.6209 | -22.1442
13.4416 | -34.6213 | -22.1526
13.4477 | -34.6218 | -22.1607
13.4538 | -34.6223 | -22.1689
13.4599 | -34.6227 | -22.1774
13.466 | -34.6232 | -22.186
13.4721 | -34.6236 | -22.194
13.4782 | -34.6241 | -22.2028
13.4843 | -34.6245 | -22.2114
13.4904 | -34.625 | -22.2197
13.4965 | -34.6254 | -22.2283
13.5026 | -34.6258 | -22.2369
13.5087 | -34.6263 | -22.2454
13.5148 | -34.6267 | -22.254
13.5209 | -34.6271 | -22.2624
13.527 | -34.6276 | -22.2708
13.5331 | -34.6281 | -22.2791
13.5393 | -34.6285 | -22.2878
13.5454 | -34.629 | -22.2962
13.5515 | -34.6294 | -22.3049
13.5576 | -34.6298 | -22.3132
13.5637 | -34.6303 | -22.3216
13.5698 | -34.6307 | -22.33
13.5759 | -34.6312 | -22.338
13.582 | -34.6317 | -22.3466
13.5881 | -34.6321 | -22.3549
13.5942 | -34.6325 | -22.3634
13.6003 | -34.633 | -22.3719
13.6064 | -34.6334 | -22.3801
13.6125 | -34.6338 | -22.3886
13.6186 | -34.6343 | -22.397
13.6247 | -34.6348 | -22.4055
13.6308 | -34.6352 | -22.4141
13.6369 | -34.6356 | -22.4225
13.643 | -34.636 | -22.4311
13.6491 | -34.6364 | -22.4394
13.6552 | -34.6369 | -22.448
13.6613 | -34.6374 | -22.4565
13.6674 | -34.6378 | -22.465
13.6735 | -34.6382 | -22.4733
13.6796 | -34.6387 | -22.4819
13.6858 | -34.6391 | -22.4897
13.6919 | -34.6396 | -22.498
13.698 | -34.64 | -22.5065
13.7041 | -34.6405 | -22.515
13.7102 | -34.6409 | -22.5234
13.7163 | -34.6414 | -22.5318
13.7224 | -34.6418 | -22.5402
13.7285 | -34.6423 | -22.5485
13.7346 | -34.6427 | -22.557
13.7407 | -34.6432 | -22.5654
13.7468 | -34.6436 | -22.5738
13.7529 | -34.6441 | -22.5819
13.759 | -34.6445 | -22.5903
13.7651 | -34.6449 | -22.5989
13.7712 | -34.6454 | -22.6073
13.7773 | -34.6458 | -22.6157
13.7834 | -34.6462 | -22.6241
13.7895 | -34.6467 | -22.6323
13.7956 | -34.6472 | -22.6408
13.8017 | -34.6476 | -22.6491
13.8078 | -34.6481 | -22.6576
13.8139 | -34.6485 | -22.666
13.82 | -34.6489 | -22.6744
13.8262 | -34.6493 | -22.6827
13.8323 | -34.6498 | -22.6913
13.8384 | -34.6503 | -22.6997
13.8445 | -34.6507 | -22.708
13.8506 | -34.6512 | -22.7164
13.8567 | -34.6516 | -22.7248
13.8628 | -34.6521 | -22.7333
13.8689 | -34.6525 | -22.7418
13.875 | -34.653 | -22.7501
13.8811 | -34.6534 | -22.7585
13.8872 | -34.6538 | -22.7669
13.8933 | -34.6542 | -22.775
13.8994 | -34.6547 | -22.7834
13.9055 | -34.6551 | -22.7918
13.9116 | -34.6555 | -22.8002
13.9177 | -34.656 | -22.8087
13.9238 | -34.6564 | -22.817
13.9299 | -34.6569 | -22.8252
13.936 | -34.6573 | -22.8337
13.9421 | -34.6578 | -22.8419
13.9482 | -34.6582 | -22.85
13.9543 | -34.6586 | -22.8585
13.9604 | -34.6591 | -22.8668
13.9665 | -34.6595 | -22.875
13.9727 | -34.66 | -22.8834
13.9788 | -34.6604 | -22.8918
13.9849 | -34.6609 | -22.9003
13.991 | -34.6613 | -22.9087
13.9971 | -34.6617 | -22.9167
14.0032 | -34.6621 | -22.9252
14.0093 | -34.6626 | -22.9335
14.0154 | -34.663 | -22.9417
14.0215 | -34.6635 | -22.9499
14.0276 | -34.664 | -22.9583
14.0337 | -34.6644 | -22.9665
14.0398 | -34.6649 | -22.9749
14.0459 | -34.6653 | -22.9832
14.052 | -34.6658 | -22.9917
14.0581 | -34.6662 | -23
14.0642 | -34.6666 | -23.0084
14.0703 | -34.6671 | -23.0167
14.0764 | -34.6675 | -23.0249
14.0825 | -34.668 | -23.0331
14.0886 | -34.6684 | -23.0413
14.0947 | -34.6689 | -23.0498
14.1008 | -34.6693 | -23.0581
14.1069 | -34.6697 | -23.0661
14.1131 | -34.6702 | -23.0746
14.1192 | -34.6706 | -23.0828
14.1253 | -34.6711 | -23.091
14.1314 | -34.6715 | -23.0993
14.1375 | -34.672 | -23.1076
14.1436 | -34.6724 | -23.1158
14.1497 | -34.6729 | -23.1241
14.1558 | -34.6734 | -23.1324
14.1619 | -34.6738 | -23.1407
14.168 | -34.6743 | -23.1489
14.1741 | -34.6747 | -23.157
14.1802 | -34.6752 | -23.1653
14.1863 | -34.6756 | -23.1735
14.1924 | -34.6761 | -23.1818
14.1985 | -34.6766 | -23.19
14.2046 | -34.677 | -23.1985
14.2107 | -34.6775 | -23.2064
14.2168 | -34.6779 | -23.2148
14.2229 | -34.6783 | -23.223
14.229 | -34.6787 | -23.2314
14.2351 | -34.6792 | -23.2398
14.2412 | -34.6796 | -23.2481
14.2473 | -34.6801 | -23.2564
14.2534 | -34.6805 | -23.2645
14.2596 | -34.681 | -23.2725
14.2657 | -34.6814 | -23.2808
14.2718 | -34.6818 | -23.2892
14.2779 | -34.6823 | -23.2976
14.284 | -34.6828 | -23.3058
14.2901 | -34.6832 | -23.3142
14.2962 | -34.6837 | -23.3223
14.3023 | -34.684 | -23.3306
14.3084 | -34.6845 | -23.3389
14.3145 | -34.6849 | -23.3472
14.3206 | -34.6854 | -23.3559
14.3267 | -34.6858 | -23.364
14.3328 | -34.6863 | -23.3722
14.3389 | -34.6867 | -23.3802
14.345 | -34.6872 | -23.3883
14.3511 | -34.6876 | -23.3967
14.3572 | -34.6881 | -23.4052
14.3633 | -34.6885 | -23.4136
14.3694 | -34.689 | -23.4218
14.3755 | -34.6894 | -23.4299
14.3816 | -34.6899 | -23.4383
14.3877 | -34.6904 | -23.4463
14.3938 | -34.6908 | -23.4546
14.4 | -34.6913 | -23.4629
14.4061 | -34.6917 | -23.4712
14.4122 | -34.6921 | -23.4793
14.4183 | -34.6926 | -23.4875
14.4244 | -34.693 | -23.4955
14.4305 | -34.6935 | -23.5037
14.4366 | -34.694 | -23.5119
14.4427 | -34.6944 | -23.52
14.4488 | -34.6948 | -23.5283
14.4549 | -34.6953 | -23.5364
14.461 | -34.6958 | -23.5446
14.4671 | -34.6962 | -23.5528
14.4732 | -34.6967 | -23.5611
14.4793 | -34.6971 | -23.5693
14.4854 | -34.6976 | -23.5776
14.4915 | -34.698 | -23.5853
14.4976 | -34.6985 | -23.5934
14.5037 | -34.6989 | -23.6016
14.5098 | -34.6994 | -23.6097
14.5159 | -34.6998 | -23.6181
14.522 | -34.7003 | -23.6261
14.5281 | -34.7008 | -23.6344
14.5342 | -34.7013 | -23.6425
14.5403 | -34.7017 | -23.6508
14.5465 | -34.7021 | -23.6588
14.5526 | -34.7026 | -23.6671
14.5587 | -34.7031 | -23.6753
14.5648 | -34.7035 | -23.6832
14.5709 | -34.704 | -23.6913
14.577 | -34.7044 | -23.6996
14.5831 | -34.7049 | -23.7078
14.5892 | -34.7053 | -23.7159
14.5953 | -34.7058 | -23.7241
14.6014 | -34.7063 | -23.7324
14.6075 | -34.7067 | -23.7406
14.6136 | -34.7071 | -23.7487
14.6197 | -34.7076 | -23.757
14.6258 | -34.708 | -23.7654
14.6319 | -34.7085 | -23.7734
14.638 | -34.7089 | -23.7816
14.6441 | -34.7094 | -23.7897
14.6502 | -34.7098 | -23.798
14.6563 | -34.7103 | -23.8061
14.6624 | -34.7107 | -23.8142
14.6685 | -34.7112 | -23.8221
14.6746 | -34.7116 | -23.8304
14.6807 | -34.7121 | -23.8385
14.6869 | -34.7126 | -23.8467
14.693 | -34.7131 | -23.8548
14.6991 | -34.7135 | -23.863
14.7052 | -34.714 | -23.871
14.7113 | -34.7145 | -23.8791
14.7174 | -34.715 | -23.8874
14.7235 | -34.7154 | -23.8953
14.7296 | -34.7159 | -23.9035
14.7357 | -34.7164 | -23.9115
14.7418 | -34.7168 | -23.9196
14.7479 | -34.7173 | -23.9277
14.754 | -34.7178 | -23.936
14.7601 | -34.7183 | -23.9442
14.7662 | -34.7187 | -23.9523
14.7723 | -34.7192 | -23.9599
14.7784 | -34.7196 | -23.9681
14.7845 | -34.72 | -23.9762
14.7906 | -34.7205 | -23.9844
14.7967 | -34.721 | -23.9925
14.8028 | -34.7214 | -24.0004
14.8089 | -34.7219 | -24.0084
14.815 | -34.7224 | -24.0166
14.8211 | -34.7228 | -24.0248
14.8272 | -34.7233 | -24.0329
14.8334 | -34.7238 | -24.041
14.8395 | -34.7242 | -24.0491
14.8456 | -34.7247 | -24.0572
14.8517 | -34.7252 | -24.0654
14.8578 | -34.7256 | -24.0735
14.8639 | -34.7261 | -24.0817
14.87 | -34.7266 | -24.0898
14.8761 | -34.727 | -24.0978
14.8822 | -34.7275 | -24.106
14.8883 | -34.728 | -24.1141
14.8944 | -34.7284 | -24.1221
14.9005 | -34.7289 | -24.1301
14.9066 | -34.7294 | -24.1382
14.9127 | -34.7299 | -24.146
14.9188 | -34.7303 | -24.1541
14.9249 | -34.7308 | -24.1623
14.931 | -34.7312 | -24.1705
14.9371 | -34.7317 | -24.1786
14.9432 | -34.7321 | -24.1866
14.9493 | -34.7326 | -24.1945
14.9554 | -34.7331 | -24.2025
14.9615 | -34.7335 | -24.2106
14.9676 | -34.734 | -24.219
14.9738 | -34.7345 | -24.227
14.9799 | -34.7349 | -24.235
14.986 | -34.7354 | -24.2428
14.9921 | -34.7358 | -24.2508
14.9982 | -34.7363 | -24.2589
15.0043 | -34.7367 | -24.2667
15.0104 | -34.7372 | -24.2749
15.0165 | -34.7377 | -24.283
15.0226 | -34.7382 | -24.2909
15.0287 | -34.7387 | -24.2989
15.0348 | -34.7392 | -24.307
15.0409 | -34.7396 | -24.315
15.047 | -34.7401 | -24.3231
15.0531 | -34.7406 | -24.3311
15.0592 | -34.7411 | -24.3392
15.0653 | -34.7415 | -24.3471
15.0714 | -34.742 | -24.355
15.0775 | -34.7425 | -24.363
15.0836 | -34.7429 | -24.3709
15.0897 | -34.7435 | -24.3789
15.0958 | -34.744 | -24.3869
15.1019 | -34.7444 | -24.395
15.108 | -34.7449 | -24.4028
15.1141 | -34.7453 | -24.411
15.1203 | -34.7458 | -24.4191
15.1264 | -34.7463 | -24.4269
15.1325 | -34.7468 | -24.4348
15.1386 | -34.7473 | -24.4426
15.1447 | -34.7477 | -24.4508
15.1508 | -34.7482 | -24.4588
15.1569 | -34.7487 | -24.4669
15.163 | -34.7492 | -24.4746
15.1691 | -34.7497 | -24.4825
15.1752 | -34.7501 | -24.4905
15.1813 | -34.7506 | -24.4985
15.1874 | -34.7511 | -24.5066
15.1935 | -34.7516 | -24.5144
15.1996 | -34.752 | -24.522
15.2057 | -34.7525 | -24.53
15.2118 | -34.753 | -24.5379
15.2179 | -34.7535 | -24.546
15.224 | -34.7539 | -24.5539
15.2301 | -34.7544 | -24.5618
15.2362 | -34.7549 | -24.5699
15.2423 | -34.7554 | -24.578
15.2484 | -34.7558 | -24.5859
15.2545 | -34.7563 | -24.594
15.2607 | -34.7568 | -24.6021
15.2668 | -34.7573 | -24.6098
15.2729 | -34.7578 | -24.6178
15.279 | -34.7583 | -24.626
15.2851 | -34.7588 | -24.6338
15.2912 | -34.7593 | -24.6418
15.2973 | -34.7597 | -24.6499
15.3034 | -34.7602 | -24.6577
15.3095 | -34.7607 | -24.6658
15.3156 | -34.7612 | -24.6737
15.3217 | -34.7616 | -24.6818
15.3278 | -34.7621 | -24.6898
15.3339 | -34.7626 | -24.698
15.34 | -34.7631 | -24.7057
15.3461 | -34.7636 | -24.7135
15.3522 | -34.7641 | -24.7215
15.3583 | -34.7645 | -24.7299
15.3644 | -34.765 | -24.7377
15.3705 | -34.7656 | -24.7453
15.3766 | -34.766 | -24.7534
15.3827 | -34.7665 | -24.7612
15.3888 | -34.767 | -24.7692
15.3949 | -34.7674 | -24.7772
15.401 | -34.7679 | -24.7852
15.4072 | -34.7684 | -24.7929
15.4133 | -34.7688 | -24.8009
15.4194 | -34.7693 | -24.8089
15.4255 | -34.7698 | -24.8171
15.4316 | -34.7703 | -24.8251
15.4377 | -34.7708 | -24.8331
15.4438 | -34.7713 | -24.8409
15.4499 | -34.7717 | -24.8487
15.456 | -34.7722 | -24.8567
15.4621 | -34.7727 | -24.8645
15.4682 | -34.7732 | -24.8723
15.4743 | -34.7737 | -24.8798
15.4804 | -34.7741 | -24.8877
15.4865 | -34.7746 | -24.8957
15.4926 | -34.7751 | -24.9038
15.4987 | -34.7756 | -24.9117
15.5048 | -34.7761 | -24.9197
15.5109 | -34.7765 | -24.9274
15.517 | -34.7771 | -24.9353
15.5231 | -34.7776 | -24.9435
15.5292 | -34.778 | -24.9514
15.5353 | -34.7785 | -24.9596
15.5414 | -34.779 | -24.9676
15.5476 | -34.7796 | -24.9752
15.5537 | -34.7801 | -24.983
15.5598 | -34.7806 | -24.991
15.5659 | -34.7811 | -24.9989
15.572 | -34.7816 | -25.0067
15.5781 | -34.7821 | -25.0145
15.5842 | -34.7826 | -25.0224
15.5903 | -34.7831 | -25.03
15.5964 | -34.7836 | -25.0377
15.6025 | -34.7841 | -25.0457
15.6086 | -34.7846 | -25.0537
15.6147 | -34.7851 | -25.0616
15.6208 | -34.7856 | -25.069
15.6269 | -34.7861 | -25.0768
15.633 | -34.7866 | -25.0844
15.6391 | -34.7871 | -25.0924
15.6452 | -34.7876 | -25.1004
15.6513 | -34.7881 | -25.1078
15.6574 | -34.7885 | -25.1157
15.6635 | -34.789 | -25.1236
15.6696 | -34.7895 | -25.1314
15.6757 | -34.79 | -25.1393
15.6818 | -34.7905 | -25.1471
15.688 | -34.791 | -25.1552
15.6941 | -34.7914 | -25.1629
15.7002 | -34.7919 | -25.1708
15.7063 | -34.7924 | -25.1787
15.7124 | -34.7929 | -25.1868
15.7185 | -34.7933 | -25.1947
15.7246 | -34.7938 | -25.2028
15.7307 | -34.7943 | -25.2105
15.7368 | -34.7948 | -25.2183
15.7429 | -34.7953 | -25.2261
15.749 | -34.7958 | -25.2339
15.7551 | -34.7963 | -25.242
15.7612 | -34.7968 | -25.2498
15.7673 | -34.7973 | -25.2575
15.7734 | -34.7978 | -25.2653
15.7795 | -34.7983 | -25.2733
15.7856 | -34.7987 | -25.2812
15.7917 | -34.7993 | -25.2891
15.7978 | -34.7997 | -25.2971
15.8039 | -34.8003 | -25.305
15.81 | -34.8007 | -25.313
15.8161 | -34.8012 | -25.3209
15.8222 | -34.8018 | -25.3288
15.8283 | -34.8023 | -25.3366
15.8345 | -34.8028 | -25.3443
15.8406 | -34.8032 | -25.3519
15.8467 | -34.8037 | -25.3597
15.8528 | -34.8042 | -25.3676
15.8589 | -34.8047 | -25.3754
15.865 | -34.8053 | -25.3829
15.8711 | -34.8058 | -25.3909
15.8772 | -34.8062 | -25.3986
15.8833 | -34.8067 | -25.4065
15.8894 | -34.8072 | -25.4145
15.8955 | -34.8077 | -25.4223
15.9016 | -34.8082 | -25.4299
15.9077 | -34.8086 | -25.4379
15.9138 | -34.8091 | -25.4459
15.9199 | -34.8096 | -25.4535
15.926 | -34.81 | -25.4613
15.9321 | -34.8105 | -25.469
15.9382 | -34.811 | -25.4766
15.9443 | -34.8115 | -25.4844
15.9504 | -34.812 | -25.4924
15.9565 | -34.8124 | -25.5001
15.9626 | -34.8129 | -25.508
15.9687 | -34.8133 | -25.5153
15.9749 | -34.8138 | -25.5228
15.981 | -34.8143 | -25.5304
15.9871 | -34.8148 | -25.5385
15.9932 | -34.8153 | -25.5462
15.9993 | -34.8158 | -25.5537
16.0054 | -34.8163 | -25.5611
16.0115 | -34.8168 | -25.5684
16.0176 | -34.8173 | -25.5764
16.0237 | -34.8178 | -25.5844
16.0298 | -34.8183 | -25.592
16.0359 | -34.8187 | -25.5997
16.042 | -34.8193 | -25.6078
16.0481 | -34.8197 | -25.6159
16.0542 | -34.8203 | -25.6238
16.0603 | -34.8207 | -25.6316
16.0664 | -34.8212 | -25.6395
16.0725 | -34.8217 | -25.6474
16.0786 | -34.8222 | -25.6547
16.0847 | -34.8227 | -25.6624
16.0908 | -34.8232 | -25.6701
16.0969 | -34.8236 | -25.6782
16.103 | -34.8241 | -25.6861
16.1091 | -34.8246 | -25.6938
16.1152 | -34.8251 | -25.7018
16.1214 | -34.8256 | -25.7096
16.1275 | -34.826 | -25.7174
16.1336 | -34.8265 | -25.7252
16.1397 | -34.8271 | -25.733
16.1458 | -34.8275 | -25.7406
16.1519 | -34.828 | -25.7487
16.158 | -34.8285 | -25.7565
16.1641 | -34.829 | -25.7646
16.1702 | -34.8295 | -25.7724
16.1763 | -34.83 | -25.7802
16.1824 | -34.8304 | -25.7875
16.1885 | -34.8309 | -25.7949
16.1946 | -34.8314 | -25.8024
16.2007 | -34.8318 | -25.8101
16.2068 | -34.8323 | -25.8178
16.2129 | -34.8328 | -25.8256
16.219 | -34.8334 | -25.833
16.2251 | -34.8338 | -25.8405
16.2312 | -34.8344 | -25.8484
16.2373 | -34.8349 | -25.8562
16.2434 | -34.8354 | -25.8642
16.2495 | -34.8358 | -25.8716
16.2556 | -34.8364 | -25.8793
16.2618 | -34.8369 | -25.8873
16.2679 | -34.8374 | -25.8954
16.274 | -34.8378 | -25.9029
16.2801 | -34.8383 | -25.9107
16.2862 | -34.8388 | -25.9185
16.2923 | -34.8393 | -25.9262
16.2984 | -34.8398 | -25.9336
16.3045 | -34.8403 | -25.9413
16.3106 | -34.8408 | -25.9495
16.3167 | -34.8412 | -25.9572
16.3228 | -34.8417 | -25.9647
16.3289 | -34.8422 | -25.9723
16.335 | -34.8427 | -25.9798
16.3411 | -34.8432 | -25.9876
16.3472 | -34.8436 | -25.9954
16.3533 | -34.8442 | -26.003
16.3594 | -34.8447 | -26.0105
16.3655 | -34.8452 | -26.0181
16.3716 | -34.8456 | -26.0256
16.3777 | -34.8461 | -26.0331
16.3838 | -34.8466 | -26.0408
16.3899 | -34.8471 | -26.0485
16.396 | -34.8476 | -26.0566
16.4021 | -34.8481 | -26.0644
16.4083 | -34.8486 | -26.072
16.4144 | -34.8491 | -26.0796
16.4205 | -34.8495 | -26.0872
16.4266 | -34.85 | -26.095
16.4327 | -34.8505 | -26.1027
16.4388 | -34.851 | -26.1106
16.4449 | -34.8515 | -26.1184
16.451 | -34.852 | -26.1264
16.4571 | -34.8525 | -26.1341
16.4632 | -34.853 | -26.1415
16.4693 | -34.8534 | -26.1493
16.4754 | -34.8539 | -26.1572
16.4815 | -34.8544 | -26.1649
16.4876 | -34.8549 | -26.1723
16.4937 | -34.8554 | -26.1799
16.4998 | -34.8559 | -26.1872
16.5059 | -34.8563 | -26.1948
16.512 | -34.8568 | -26.2029
16.5181 | -34.8572 | -26.2107
16.5242 | -34.8577 | -26.2184
16.5303 | -34.8581 | -26.2261
16.5364 | -34.8586 | -26.2338
16.5425 | -34.8591 | -26.2413
16.5487 | -34.8596 | -26.2489
16.5548 | -34.86 | -26.2565
16.5609 | -34.8605 | -26.2644
16.567 | -34.861 | -26.2722
16.5731 | -34.8615 | -26.2798
16.5792 | -34.862 | -26.2875
16.5853 | -34.8625 | -26.2949
16.5914 | -34.863 | -26.3028
16.5975 | -34.8634 | -26.3106
16.6036 | -34.8638 | -26.3181
16.6097 | -34.8643 | -26.3258
16.6158 | -34.8648 | -26.3334
16.6219 | -34.8653 | -26.341
16.628 | -34.8658 | -26.3484
16.6341 | -34.8663 | -26.3564
16.6402 | -34.8668 | -26.3643
16.6463 | -34.8672 | -26.3719
16.6524 | -34.8677 | -26.3794
16.6585 | -34.8681 | -26.3869
16.6646 | -34.8686 | -26.3946
16.6707 | -34.869 | -26.4018
16.6768 | -34.8695 | -26.4095
16.6829 | -34.87 | -26.4169
16.689 | -34.8705 | -26.4247
16.6952 | -34.8709 | -26.4323
16.7013 | -34.8714 | -26.4398
16.7074 | -34.8719 | -26.4471
16.7135 | -34.8724 | -26.4549
16.7196 | -34.8729 | -26.4624
16.7257 | -34.8734 | -26.4699
16.7318 | -34.8738 | -26.4776
16.7379 | -34.8743 | -26.4853
16.744 | -34.8748 | -26.4925
16.7501 | -34.8753 | -26.5001
16.7562 | -34.8758 | -26.5077
16.7623 | -34.8763 | -26.5153
16.7684 | -34.8767 | -26.523
16.7745 | -34.8772 | -26.5306
16.7806 | -34.8777 | -26.538
16.7867 | -34.8781 | -26.5454
16.7928 | -34.8786 | -26.5527
16.7989 | -34.8791 | -26.5604
16.805 | -34.8796 | -26.5679
16.8111 | -34.88 | -26.5753
16.8172 | -34.8804 | -26.5834
16.8233 | -34.8809 | -26.5913
16.8294 | -34.8814 | -26.599
16.8356 | -34.8819 | -26.6069
16.8417 | -34.8823 | -26.6145
16.8478 | -34.8828 | -26.622
16.8539 | -34.8833 | -26.6295
16.86 | -34.8837 | -26.6373
16.8661 | -34.8842 | -26.6449
16.8722 | -34.8846 | -26.6531
16.8783 | -34.8851 | -26.6608
16.8844 | -34.8857 | -26.6682
16.8905 | -34.8862 | -26.6756
16.8966 | -34.8866 | -26.6833
16.9027 | -34.8871 | -26.6911
16.9088 | -34.8876 | -26.6988
16.9149 | -34.8881 | -26.7061
16.921 | -34.8886 | -26.7136
16.9271 | -34.8891 | -26.7211
16.9332 | -34.8896 | -26.7289
16.9393 | -34.8901 | -26.7364
16.9454 | -34.8906 | -26.7437
16.9515 | -34.8911 | -26.7515
16.9576 | -34.8915 | -26.759
16.9637 | -34.892 | -26.7665
16.9698 | -34.8925 | -26.7741
16.9759 | -34.893 | -26.7818
16.9821 | -34.8935 | -26.7897
16.9882 | -34.8939 | -26.7969
16.9943 | -34.8944 | -26.8046
17.0004 | -34.8949 | -26.8121
17.0065 | -34.8953 | -26.8195
17.0126 | -34.8958 | -26.8272
17.0187 | -34.8963 | -26.8349
17.0248 | -34.8968 | -26.8427
17.0309 | -34.8973 | -26.8501
17.037 | -34.8978 | -26.8577
17.0431 | -34.8982 | -26.8655
17.0492 | -34.8986 | -26.8729
17.0553 | -34.8991 | -26.8801
17.0614 | -34.8996 | -26.8875
17.0675 | -34.9 | -26.8949
17.0736 | -34.9005 | -26.9026
17.0797 | -34.901 | -26.91
17.0858 | -34.9014 | -26.9174
17.0919 | -34.902 | -26.9246
17.098 | -34.9025 | -26.932
17.1041 | -34.903 | -26.9394
17.1102 | -34.9035 | -26.9469
17.1163 | -34.904 | -26.9544
17.1225 | -34.9045 | -26.9617
17.1286 | -34.9049 | -26.9691
17.1347 | -34.9054 | -26.9766
17.1408 | -34.9059 | -26.9845
17.1469 | -34.9064 | -26.9921
17.153 | -34.9069 | -26.9996
17.1591 | -34.9073 | -27.0067
17.1652 | -34.9078 | -27.0143
17.1713 | -34.9083 | -27.0216
17.1774 | -34.9088 | -27.029
17.1835 | -34.9092 | -27.0364
17.1896 | -34.9097 | -27.0439
17.1957 | -34.9102 | -27.0513
17.2018 | -34.9107 | -27.0589
17.2079 | -34.9112 | -27.0664
17.214 | -34.9116 | -27.074
17.2201 | -34.9121 | -27.0813
17.2262 | -34.9126 | -27.0891
17.2323 | -34.9131 | -27.0963
17.2384 | -34.9135 | -27.104
17.2445 | -34.9139 | -27.1115
17.2506 | -34.9144 | -27.1188
17.2567 | -34.9148 | -27.1266
17.2628 | -34.9153 | -27.1341
17.269 | -34.9157 | -27.1415
17.2751 | -34.9162 | -27.1492
17.2812 | -34.9167 | -27.1566
17.2873 | -34.9171 | -27.1642
17.2934 | -34.9175 | -27.1716
17.2995 | -34.918 | -27.1792
17.3056 | -34.9185 | -27.1867
17.3117 | -34.919 | -27.1944
17.3178 | -34.9194 | -27.202
17.3239 | -34.9199 | -27.2095
17.33 | -34.9203 | -27.2171
17.3361 | -34.9208 | -27.2245
17.3422 | -34.9211 | -27.2318
17.3483 | -34.9216 | -27.2394
17.3544 | -34.9221 | -27.2471
17.3605 | -34.9225 | -27.2547
17.3666 | -34.923 | -27.2624
17.3727 | -34.9235 | -27.2696
17.3788 | -34.9239 | -27.2766
17.3849 | -34.9243 | -27.2841
17.391 | -34.9248 | -27.2916
17.3971 | -34.9252 | -27.2991
17.4032 | -34.9257 | -27.3066
17.4094 | -34.9262 | -27.3142
17.4155 | -34.9266 | -27.3213
17.4216 | -34.9271 | -27.3288
17.4277 | -34.9276 | -27.3366
17.4338 | -34.928 | -27.3442
17.4399 | -34.9285 | -27.352
17.446 | -34.9289 | -27.359
17.4521 | -34.9294 | -27.3667
17.4582 | -34.9298 | -27.3738
17.4643 | -34.9303 | -27.3811
17.4704 | -34.9308 | -27.3888
17.4765 | -34.9313 | -27.3962
17.4826 | -34.9317 | -27.4034
17.4887 | -34.9322 | -27.4111
17.4948 | -34.9326 | -27.4188
17.5009 | -34.9331 | -27.4262
17.507 | -34.9336 | -27.4337
17.5131 | -34.934 | -27.4412
17.5192 | -34.9344 | -27.4484
17.5253 | -34.9349 | -27.4561
17.5314 | -34.9354 | -27.4635
17.5375 | -34.9358 | -27.471
17.5436 | -34.9363 | -27.4787
17.5497 | -34.9368 | -27.4859
17.5559 | -34.9372 | -27.4932
17.562 | -34.9377 | -27.5004
17.5681 | -34.9381 | -27.5078
17.5742 | -34.9386 | -27.5152
17.5803 | -34.9391 | -27.5223
17.5864 | -34.9395 | -27.5297
17.5925 | -34.94 | -27.5373
17.5986 | -34.9405 | -27.5446
17.6047 | -34.941 | -27.5523
17.6108 | -34.9414 | -27.5596
17.6169 | -34.9418 | -27.5672
17.623 | -34.9423 | -27.5747
17.6291 | -34.9428 | -27.582
17.6352 | -34.9432 | -27.5893
17.6413 | -34.9437 | -27.597
17.6474 | -34.9442 | -27.604
17.6535 | -34.9446 | -27.6113
17.6596 | -34.945 | -27.6185
17.6657 | -34.9455 | -27.6258
17.6718 | -34.946 | -27.6332
17.6779 | -34.9464 | -27.6406
17.684 | -34.9469 | -27.6479
17.6901 | -34.9473 | -27.6555
17.6963 | -34.9478 | -27.6627
17.7024 | -34.9483 | -27.67
17.7085 | -34.9487 | -27.6771
17.7146 | -34.9491 | -27.6843
17.7207 | -34.9496 | -27.6915
17.7268 | -34.9501 | -27.6987
17.7329 | -34.9505 | -27.706
17.739 | -34.9509 | -27.7136
17.7451 | -34.9513 | -27.7207
17.7512 | -34.9518 | -27.7284
17.7573 | -34.9523 | -27.736
17.7634 | -34.9527 | -27.7429
17.7695 | -34.9532 | -27.7503
17.7756 | -34.9536 | -27.7579
17.7817 | -34.9541 | -27.7653
17.7878 | -34.9545 | -27.7726
17.7939 | -34.955 | -27.7798
17.8 | -34.9555 | -27.7868
17.8061 | -34.9559 | -27.7942
17.8122 | -34.9564 | -27.8014
17.8183 | -34.9569 | -27.8088
17.8244 | -34.9574 | -27.8161
17.8305 | -34.9578 | -27.8235
17.8366 | -34.9583 | -27.8301
17.8428 | -34.9588 | -27.8375
17.8489 | -34.9593 | -27.8446
17.855 | -34.9597 | -27.8518
17.8611 | -34.9602 | -27.8592
17.8672 | -34.9607 | -27.8661
17.8733 | -34.9611 | -27.8735
17.8794 | -34.9616 | -27.8808
17.8855 | -34.9621 | -27.8882
17.8916 | -34.9625 | -27.8957
17.8977 | -34.9629 | -27.903
17.9038 | -34.9635 | -27.9102
17.9099 | -34.9639 | -27.9174
17.916 | -34.9644 | -27.9248
17.9221 | -34.9649 | -27.9321
17.9282 | -34.9653 | -27.9394
17.9343 | -34.9658 | -27.9465
17.9404 | -34.9664 | -27.9537
17.9465 | -34.9668 | -27.9609
17.9526 | -34.9673 | -27.968
17.9587 | -34.9677 | -27.9755
17.9648 | -34.9682 | -27.9827
17.9709 | -34.9686 | -27.99
17.977 | -34.9691 | -27.9972
17.9832 | -34.9696 | -28.0048
17.9893 | -34.97 | -28.0122
17.9954 | -34.9705 | -28.02
18.0015 | -34.9709 | -28.0272
18.0076 | -34.9714 | -28.034
18.0137 | -34.9718 | -28.0415
18.0198 | -34.9722 | -28.0486
18.0259 | -34.9727 | -28.0559
18.032 | -34.9731 | -28.063
18.0381 | -34.9735 | -28.0702
18.0442 | -34.974 | -28.0771
18.0503 | -34.9744 | -28.0844
18.0564 | -34.9749 | -28.0912
18.0625 | -34.9754 | -28.0983
18.0686 | -34.9758 | -28.1054
18.0747 | -34.9763 | -28.1128
18.0808 | -34.9767 | -28.1201
18.0869 | -34.9772 | -28.1272
18.093 | -34.9777 | -28.134
18.0991 | -34.9782 | -28.1413
18.1052 | -34.9786 | -28.1488
18.1113 | -34.9791 | -28.1557
18.1174 | -34.9796 | -28.1628
18.1236 | -34.98 | -28.1701
18.1297 | -34.9805 | -28.1776
18.1358 | -34.981 | -28.1852
18.1419 | -34.9814 | -28.1924
18.148 | -34.9818 | -28.1994
18.1541 | -34.9824 | -28.2065
18.1602 | -34.9828 | -28.214
18.1663 | -34.9833 | -28.221
18.1724 | -34.9837 | -28.2282
18.1785 | -34.9842 | -28.2352
18.1846 | -34.9846 | -28.2423
18.1907 | -34.9851 | -28.2498
18.1968 | -34.9855 | -28.2571
18.2029 | -34.9859 | -28.2644
18.209 | -34.9864 | -28.2718
18.2151 | -34.9869 | -28.2792
18.2212 | -34.9873 | -28.2862
18.2273 | -34.9878 | -28.2934
18.2334 | -34.9883 | -28.3002
18.2395 | -34.9888 | -28.3071
18.2456 | -34.9893 | -28.3145
18.2517 | -34.9897 | -28.3217
18.2578 | -34.9902 | -28.3289
18.2639 | -34.9907 | -28.3361
18.2701 | -34.9911 | -28.3434
18.2762 | -34.9915 | -28.3505
18.2823 | -34.992 | -28.3576
18.2884 | -34.9924 | -28.365
18.2945 | -34.9928 | -28.3721
18.3006 | -34.9933 | -28.3793
18.3067 | -34.9937 | -28.3863
18.3128 | -34.9942 | -28.3936
18.3189 | -34.9946 | -28.4008
18.325 | -34.9951 | -28.4073
18.3311 | -34.9956 | -28.4143
18.3372 | -34.996 | -28.4217
18.3433 | -34.9965 | -28.4292
18.3494 | -34.997 | -28.4364
18.3555 | -34.9974 | -28.4436
18.3616 | -34.9979 | -28.4506
18.3677 | -34.9983 | -28.4581
18.3738 | -34.9987 | -28.4646
18.3799 | -34.9992 | -28.4718
18.386 | -34.9996 | -28.4786
18.3921 | -35.0001 | -28.4855
18.3982 | -35.0006 | -28.4922
18.4043 | -35.0011 | -28.4992
18.4105 | -35.0015 | -28.5058
18.4166 | -35.002 | -28.5127
18.4227 | -35.0025 | -28.5196
18.4288 | -35.0029 | -28.5265
18.4349 | -35.0033 | -28.5331
18.441 | -35.0038 | -28.5397
18.4471 | -35.0042 | -28.5468
18.4532 | -35.0047 | -28.554
18.4593 | -35.0051 | -28.561
18.4654 | -35.0056 | -28.5677
18.4715 | -35.0061 | -28.5751
18.4776 | -35.0065 | -28.5823
18.4837 | -35.007 | -28.5894
18.4898 | -35.0074 | -28.5967
18.4959 | -35.0078 | -28.604
18.502 | -35.0082 | -28.6114
18.5081 | -35.0087 | -28.6184
18.5142 | -35.009 | -28.6256
18.5203 | -35.0095 | -28.6327
18.5264 | -35.01 | -28.6399
18.5325 | -35.0104 | -28.6473
18.5386 | -35.0109 | -28.6543
18.5447 | -35.0113 | -28.6613
18.5508 | -35.0118 | -28.6679
18.557 | -35.0123 | -28.675
18.5631 | -35.0127 | -28.6823
18.5692 | -35.0131 | -28.6897
18.5753 | -35.0137 | -28.6966
18.5814 | -35.0142 | -28.7037
18.5875 | -35.0146 | -28.7105
18.5936 | -35.0151 | -28.7172
18.5997 | -35.0156 | -28.724
18.6058 | -35.0161 | -28.7304
18.6119 | -35.0166 | -28.7373
18.618 | -35.017 | -28.7446
18.6241 | -35.0174 | -28.7517
18.6302 | -35.0179 | -28.7589
18.6363 | -35.0184 | -28.7656
18.6424 | -35.0189 | -28.7723
18.6485 | -35.0194 | -28.7798
18.6546 | -35.0199 | -28.7864
18.6607 | -35.0203 | -28.7937
18.6668 | -35.0208 | -28.8008
18.6729 | -35.0213 | -28.8083
18.679 | -35.0218 | -28.8153
18.6851 | -35.0222 | -28.822
18.6912 | -35.0226 | -28.8287
18.6974 | -35.0231 | -28.8355
18.7035 | -35.0235 | -28.8428
18.7096 | -35.0241 | -28.8499
18.7157 | -35.0246 | -28.8569
18.7218 | -35.0251 | -28.8639
18.7279 | -35.0255 | -28.8705
18.734 | -35.026 | -28.8779
18.7401 | -35.0265 | -28.8846
18.7462 | -35.0269 | -28.8914
18.7523 | -35.0274 | -28.899
18.7584 | -35.0278 | -28.9061
18.7645 | -35.0283 | -28.9131
18.7706 | -35.0288 | -28.9203
18.7767 | -35.0293 | -28.9273
18.7828 | -35.0297 | -28.9339
18.7889 | -35.0302 | -28.9407
18.795 | -35.0307 | -28.9482
18.8011 | -35.0311 | -28.9551
18.8072 | -35.0316 | -28.9625
18.8133 | -35.032 | -28.9694
18.8194 | -35.0326 | -28.976
18.8255 | -35.0331 | -28.983
18.8316 | -35.0335 | -28.99
18.8377 | -35.034 | -28.9973
18.8439 | -35.0345 | -29.0042
18.85 | -35.0349 | -29.0113
18.8561 | -35.0355 | -29.0176
18.8622 | -35.0359 | -29.0246
18.8683 | -35.0364 | -29.0316
18.8744 | -35.0368 | -29.0381
18.8805 | -35.0373 | -29.0452
18.8866 | -35.0378 | -29.0518
18.8927 | -35.0382 | -29.0585
18.8988 | -35.0385 | -29.0656
18.9049 | -35.039 | -29.0725
18.911 | -35.0394 | -29.08
18.9171 | -35.0399 | -29.0873
18.9232 | -35.0403 | -29.0943
18.9293 | -35.0408 | -29.1013
18.9354 | -35.0413 | -29.1086
18.9415 | -35.0418 | -29.1156
18.9476 | -35.0422 | -29.1222
18.9537 | -35.0427 | -29.1291
18.9598 | -35.0431 | -29.1362
18.9659 | -35.0435 | -29.1434
18.972 | -35.044 | -29.15
18.9781 | -35.0445 | -29.1571
18.9843 | -35.045 | -29.164
18.9904 | -35.0455 | -29.1713
18.9965 | -35.0461 | -29.1778
19.0026 | -35.0466 | -29.1846
19.0087 | -35.047 | -29.1914
19.0148 | -35.0476 | -29.1984
19.0209 | -35.0481 | -29.2052
19.027 | -35.0486 | -29.2125
19.0331 | -35.049 | -29.2186
19.0392 | -35.0495 | -29.2258
19.0453 | -35.0499 | -29.2331
19.0514 | -35.0503 | -29.2394
19.0575 | -35.0507 | -29.2463
19.0636 | -35.0512 | -29.253
19.0697 | -35.0516 | -29.2604
19.0758 | -35.0521 | -29.2671
19.0819 | -35.0525 | -29.2744
19.088 | -35.053 | -29.2816
19.0941 | -35.0534 | -29.2889
19.1002 | -35.0538 | -29.2949
19.1063 | -35.0542 | -29.3015
19.1124 | -35.0547 | -29.3087
19.1185 | -35.0551 | -29.3155
19.1246 | -35.0555 | -29.3228
19.1308 | -35.0559 | -29.33
19.1369 | -35.0563 | -29.3369
19.143 | -35.0568 | -29.3442
19.1491 | -35.0572 | -29.3512
19.1552 | -35.0576 | -29.3578
19.1613 | -35.0581 | -29.365
19.1674 | -35.0585 | -29.3721
19.1735 | -35.059 | -29.3795
19.1796 | -35.0595 | -29.3865
19.1857 | -35.0599 | -29.3934
19.1918 | -35.0604 | -29.4006
19.1979 | -35.0609 | -29.408
19.204 | -35.0615 | -29.4149
19.2101 | -35.062 | -29.422
19.2162 | -35.0624 | -29.4289
19.2223 | -35.0629 | -29.436
19.2284 | -35.0634 | -29.4432
19.2345 | -35.0639 | -29.45
19.2406 | -35.0643 | -29.457
19.2467 | -35.0648 | -29.4636
19.2528 | -35.0651 | -29.4708
19.2589 | -35.0656 | -29.4776
19.265 | -35.066 | -29.4848
19.2712 | -35.0664 | -29.4915
19.2773 | -35.0668 | -29.4991
19.2834 | -35.0673 | -29.5056
19.2895 | -35.0678 | -29.5127
19.2956 | -35.0682 | -29.5198
19.3017 | -35.0687 | -29.5264
19.3078 | -35.0691 | -29.5339
19.3139 | -35.0695 | -29.5406
19.32 | -35.0701 | -29.5478
19.3261 | -35.0705 | -29.5547
19.3322 | -35.071 | -29.5619
19.3383 | -35.0714 | -29.5688
19.3444 | -35.0719 | -29.5757
19.3505 | -35.0724 | -29.5827
19.3566 | -35.0728 | -29.5892
19.3627 | -35.0734 | -29.5962
19.3688 | -35.0739 | -29.6028
19.3749 | -35.0744 | -29.6094
19.381 | -35.075 | -29.6168
19.3871 | -35.0754 | -29.6235
19.3932 | -35.0758 | -29.6302
19.3993 | -35.0763 | -29.6376
19.4054 | -35.0767 | -29.6448
19.4115 | -35.0772 | -29.6518
19.4177 | -35.0777 | -29.6585
19.4238 | -35.0781 | -29.6655
19.4299 | -35.0786 | -29.6722
19.436 | -35.079 | -29.6787
19.4421 | -35.0794 | -29.6853
19.4482 | -35.0798 | -29.6922
19.4543 | -35.0804 | -29.6986
19.4604 | -35.0809 | -29.7055
19.4665 | -35.0814 | -29.7122
19.4726 | -35.0818 | -29.719
19.4787 | -35.0823 | -29.7258
19.4848 | -35.0828 | -29.7327
19.4909 | -35.0832 | -29.7396
19.497 | -35.0837 | -29.7464
19.5031 | -35.0842 | -29.753
19.5092 | -35.0847 | -29.76
19.5153 | -35.0851 | -29.7668
19.5214 | -35.0856 | -29.7736
19.5275 | -35.086 | -29.7809
19.5336 | -35.0864 | -29.7877
19.5397 | -35.087 | -29.7943
19.5458 | -35.0874 | -29.8017
19.5519 | -35.0879 | -29.8086
19.5581 | -35.0884 | -29.8152
19.5642 | -35.0888 | -29.8215
19.5703 | -35.0892 | -29.8284
19.5764 | -35.0898 | -29.8352
19.5825 | -35.0903 | -29.8419
19.5886 | -35.0907 | -29.8489
19.5947 | -35.0912 | -29.8556
19.6008 | -35.0917 | -29.8626
19.6069 | -35.0921 | -29.8696
19.613 | -35.0927 | -29.8764
19.6191 | -35.0931 | -29.8839
19.6252 | -35.0936 | -29.8909
19.6313 | -35.094 | -29.8978
19.6374 | -35.0945 | -29.9048
19.6435 | -35.0949 | -29.9114
19.6496 | -35.0954 | -29.9181
19.6557 | -35.0958 | -29.9248
19.6618 | -35.0961 | -29.9314
19.6679 | -35.0965 | -29.9379
19.674 | -35.0969 | -29.9438
19.6801 | -35.0974 | -29.9508
19.6862 | -35.0979 | -29.9578
19.6923 | -35.0984 | -29.9642
19.6984 | -35.0989 | -29.9702
19.7046 | -35.0994 | -29.977
19.7107 | -35.0999 | -29.983
19.7168 | -35.1005 | -29.9902
19.7229 | -35.1008 | -29.9973
19.729 | -35.1013 | -30.0045
19.7351 | -35.1017 | -30.0112
19.7412 | -35.1022 | -30.0184
19.7473 | -35.1026 | -30.0258
19.7534 | -35.1031 | -30.0323
19.7595 | -35.1035 | -30.0394
19.7656 | -35.1039 | -30.0466
19.7717 | -35.1045 | -30.0533
19.7778 | -35.1049 | -30.0602
19.7839 | -35.1054 | -30.067
19.79 | -35.1059 | -30.0735
19.7961 | -35.1062 | -30.0801
19.8022 | -35.1068 | -30.0866
19.8083 | -35.1074 | -30.0931
19.8144 | -35.1078 | -30.1001
19.8205 | -35.1083 | -30.1065
19.8266 | -35.1087 | -30.113
19.8327 | -35.1092 | -30.1195
19.8388 | -35.1096 | -30.1262
19.845 | -35.1101 | -30.133
19.8511 | -35.1106 | -30.1399
19.8572 | -35.1111 | -30.1467
19.8633 | -35.1116 | -30.1531
19.8694 | -35.112 | -30.1598
19.8755 | -35.1125 | -30.1661
19.8816 | -35.113 | -30.1732
19.8877 | -35.1134 | -30.1803
19.8938 | -35.1139 | -30.1869
19.8999 | -35.1144 | -30.1933
19.906 | -35.1147 | -30.2001
19.9121 | -35.1152 | -30.2074
19.9182 | -35.1157 | -30.2138
19.9243 | -35.1162 | -30.2204
19.9304 | -35.1166 | -30.2275
19.9365 | -35.1171 | -30.234
19.9426 | -35.1176 | -30.2409
19.9487 | -35.118 | -30.2476
19.9548 | -35.1185 | -30.2542
19.9609 | -35.119 | -30.2607
19.967 | -35.1193 | -30.2665
19.9731 | -35.1198 | -30.2728
19.9792 | -35.1203 | -30.2794
19.9853 | -35.1208 | -30.2858
19.9915 | -35.1214 | -30.2922
19.9976 | -35.1218 | -30.2992
20.0037 | -35.1223 | -30.3063
20.0098 | -35.1227 | -30.3129
20.0159 | -35.1232 | -30.3195
20.022 | -35.1237 | -30.3273
20.0281 | -35.1241 | -30.334
20.0342 | -35.1246 | -30.3409
20.0403 | -35.1251 | -30.3481
20.0464 | -35.1255 | -30.3553
20.0525 | -35.1261 | -30.3619
20.0586 | -35.1264 | -30.3685
20.0647 | -35.1269 | -30.3752
20.0708 | -35.1273 | -30.3816
20.0769 | -35.1279 | -30.3885
20.083 | -35.1282 | -30.3959
20.0891 | -35.1287 | -30.4026
20.0952 | -35.1291 | -30.4094
20.1013 | -35.1295 | -30.4166
20.1074 | -35.13 | -30.4232
20.1135 | -35.1304 | -30.4301
20.1196 | -35.1308 | -30.4361
20.1257 | -35.1314 | -30.4432
20.1319 | -35.1319 | -30.4494
20.138 | -35.1324 | -30.4563
20.1441 | -35.1328 | -30.4631
20.1502 | -35.1333 | -30.4705
20.1563 | -35.1337 | -30.4769
20.1624 | -35.1342 | -30.4841
20.1685 | -35.1347 | -30.4901
20.1746 | -35.1351 | -30.4968
20.1807 | -35.1357 | -30.5031
20.1868 | -35.1361 | -30.5093
20.1929 | -35.1365 | -30.5157
20.199 | -35.1368 | -30.5214
20.2051 | -35.1374 | -30.5281
20.2112 | -35.1379 | -30.5347
20.2173 | -35.1383 | -30.5415
20.2234 | -35.1387 | -30.5481
20.2295 | -35.1391 | -30.5546
20.2356 | -35.1396 | -30.5611
20.2417 | -35.1401 | -30.5671
20.2478 | -35.1405 | -30.573
20.2539 | -35.1411 | -30.5803
20.26 | -35.1414 | -30.5874
20.2661 | -35.1418 | -30.5942
20.2723 | -35.1423 | -30.601
20.2784 | -35.1428 | -30.6073
20.2845 | -35.1433 | -30.6142
20.2906 | -35.1437 | -30.6205
20.2967 | -35.1441 | -30.6267
20.3028 | -35.1446 | -30.6331
20.3089 | -35.1451 | -30.6401
20.315 | -35.1455 | -30.6459
20.3211 | -35.1461 | -30.6524
20.3272 | -35.1466 | -30.6597
20.3333 | -35.147 | -30.665
20.3394 | -35.1475 | -30.6717
20.3455 | -35.148 | -30.6783
20.3516 | -35.1484 | -30.6844
20.3577 | -35.1488 | -30.6913
20.3638 | -35.1493 | -30.6985
20.3699 | -35.1497 | -30.7036
20.376 | -35.1501 | -30.7101
20.3821 | -35.1505 | -30.7174
20.3882 | -35.1509 | -30.7243
20.3943 | -35.1513 | -30.7306
20.4004 | -35.1519 | -30.7367
20.4065 | -35.1523 | -30.7427
20.4126 | -35.1528 | -30.749
20.4188 | -35.1531 | -30.7561
20.4249 | -35.1535 | -30.7627
20.431 | -35.154 | -30.7695
20.4371 | -35.1545 | -30.7766
20.4432 | -35.155 | -30.7837
20.4493 | -35.1554 | -30.7908
20.4554 | -35.1559 | -30.7969
20.4615 | -35.1563 | -30.8043
20.4676 | -35.1567 | -30.8108
20.4737 | -35.1571 | -30.8169
20.4798 | -35.1576 | -30.823
20.4859 | -35.1581 | -30.8291
20.492 | -35.1585 | -30.8359
20.4981 | -35.159 | -30.8426
20.5042 | -35.1594 | -30.8497
20.5103 | -35.1598 | -30.8557
20.5164 | -35.1603 | -30.8626
20.5225 | -35.1606 | -30.869
20.5286 | -35.1611 | -30.876
20.5347 | -35.1615 | -30.8828
20.5408 | -35.1619 | -30.8888
20.5469 | -35.1623 | -30.8959
20.553 | -35.1628 | -30.9023
20.5592 | -35.1633 | -30.9094
20.5653 | -35.1638 | -30.9154
20.5714 | -35.1642 | -30.9229
20.5775 | -35.1648 | -30.9294
20.5836 | -35.1652 | -30.9353
20.5897 | -35.1657 | -30.9417
20.5958 | -35.1661 | -30.9491
20.6019 | -35.1665 | -30.955
20.608 | -35.167 | -30.9619
20.6141 | -35.1674 | -30.9686
20.6202 | -35.1679 | -30.9754
20.6263 | -35.1684 | -30.9827
20.6324 | -35.1688 | -30.9902
20.6385 | -35.1691 | -30.9972
20.6446 | -35.1696 | -31.0032
20.6507 | -35.1699 | -31.0092
20.6568 | -35.1705 | -31.0163
20.6629 | -35.171 | -31.0226
20.669 | -35.1714 | -31.0298
20.6751 | -35.1718 | -31.0366
20.6812 | -35.1721 | -31.0425
20.6873 | -35.1725 | -31.049
20.6934 | -35.1729 | -31.0555
20.6995 | -35.1733 | -31.0621
20.7057 | -35.1738 | -31.0686
20.7118 | -35.1742 | -31.0754
20.7179 | -35.1748 | -31.0822
20.724 | -35.1752 | -31.088
20.7301 | -35.1757 | -31.0945
20.7362 | -35.1761 | -31.1009
20.7423 | -35.1767 | -31.1075
20.7484 | -35.1772 | -31.1136
20.7545 | -35.1776 | -31.1192
20.7606 | -35.1781 | -31.1265
20.7667 | -35.1786 | -31.1333
20.7728 | -35.179 | -31.1399
20.7789 | -35.1795 | -31.1461
20.785 | -35.18 | -31.1531
20.7911 | -35.1803 | -31.1594
20.7972 | -35.1807 | -31.1655
20.8033 | -35.1813 | -31.1723
20.8094 | -35.1817 | -31.1795
20.8155 | -35.1822 | -31.1868
20.8216 | -35.1827 | -31.1935
20.8277 | -35.1831 | -31.1997
20.8338 | -35.1836 | -31.2064
20.8399 | -35.1841 | -31.2134
20.8461 | -35.1847 | -31.2198
20.8522 | -35.185 | -31.2269
20.8583 | -35.1854 | -31.2334
20.8644 | -35.1859 | -31.2401
20.8705 | -35.1864 | -31.2478
20.8766 | -35.1868 | -31.2543
20.8827 | -35.1872 | -31.2616
20.8888 | -35.1876 | -31.2683
20.8949 | -35.1879 | -31.2735
20.901 | -35.1884 | -31.28
20.9071 | -35.1889 | -31.2865
20.9132 | -35.1893 | -31.293
20.9193 | -35.1897 | -31.2989
20.9254 | -35.1902 | -31.3055
20.9315 | -35.1906 | -31.3119
20.9376 | -35.191 | -31.3184
20.9437 | -35.1913 | -31.3249
20.9498 | -35.1918 | -31.3311
20.9559 | -35.1922 | -31.3374
20.962 | -35.1927 | -31.3439
20.9681 | -35.1931 | -31.3496
20.9742 | -35.1936 | -31.3561
20.9803 | -35.1939 | -31.3624
20.9864 | -35.1942 | -31.3694
20.9926 | -35.1947 | -31.3756
20.9987 | -35.1949 | -31.3818
21.0048 | -35.1954 | -31.3886
21.0109 | -35.1959 | -31.3951
21.017 | -35.1963 | -31.4018
21.0231 | -35.1968 | -31.4078
21.0292 | -35.1973 | -31.4146
21.0353 | -35.1977 | -31.4214
21.0414 | -35.1982 | -31.4284
21.0475 | -35.1985 | -31.4354
21.0536 | -35.1989 | -31.4415
21.0597 | -35.1993 | -31.4482
21.0658 | -35.1998 | -31.4542
21.0719 | -35.2001 | -31.4609
21.078 | -35.2007 | -31.4675
21.0841 | -35.2011 | -31.474
21.0902 | -35.2017 | -31.4799
21.0963 | -35.2022 | -31.4864
21.1024 | -35.2026 | -31.493
21.1085 | -35.2031 | -31.4992
21.1146 | -35.2035 | -31.5058
21.1207 | -35.2039 | -31.5128
21.1268 | -35.2043 | -31.5192
21.133 | -35.2048 | -31.5266
21.1391 | -35.2052 | -31.5337
21.1452 | -35.2055 | -31.5397
21.1513 | -35.2061 | -31.5463
21.1574 | -35.2066 | -31.5531
21.1635 | -35.2071 | -31.5588
21.1696 | -35.2076 | -31.5646
21.1757 | -35.208 | -31.5706
21.1818 | -35.2083 | -31.5771
21.1879 | -35.2087 | -31.5839
21.194 | -35.2092 | -31.5904
21.2001 | -35.2096 | -31.5971
21.2062 | -35.21 | -31.6041
21.2123 | -35.2104 | -31.6101
21.2184 | -35.2108 | -31.6165
21.2245 | -35.2113 | -31.6228
21.2306 | -35.2117 | -31.6291
21.2367 | -35.2123 | -31.6358
21.2428 | -35.2127 | -31.6421
21.2489 | -35.213 | -31.6474
21.255 | -35.2135 | -31.6545
21.2611 | -35.214 | -31.661
21.2672 | -35.2144 | -31.6672
21.2733 | -35.2148 | -31.6736
21.2795 | -35.2153 | -31.6805
21.2856 | -35.2157 | -31.6874
21.2917 | -35.2162 | -31.6931
21.2978 | -35.2166 | -31.6995
21.3039 | -35.2171 | -31.7059
21.31 | -35.2177 | -31.7122
21.3161 | -35.2183 | -31.7186
21.3222 | -35.2186 | -31.7245
21.3283 | -35.219 | -31.7303
21.3344 | -35.2194 | -31.7375
21.3405 | -35.22 | -31.7443
21.3466 | -35.2204 | -31.7503
21.3527 | -35.2209 | -31.7562
21.3588 | -35.2213 | -31.7632
21.3649 | -35.2218 | -31.7698
21.371 | -35.2222 | -31.7758
21.3771 | -35.2228 | -31.7822
21.3832 | -35.2233 | -31.7892
21.3893 | -35.2237 | -31.7955
21.3954 | -35.2242 | -31.8016
21.4015 | -35.2247 | -31.8078
21.4076 | -35.2251 | -31.8142
21.4137 | -35.2255 | -31.8205
21.4199 | -35.226 | -31.8262
21.426 | -35.2264 | -31.8335
21.4321 | -35.2269 | -31.8403
21.4382 | -35.2275 | -31.8465
21.4443 | -35.2279 | -31.8526
21.4504 | -35.2285 | -31.8592
21.4565 | -35.229 | -31.8656
21.4626 | -35.2294 | -31.8716
21.4687 | -35.2299 | -31.8778
21.4748 | -35.2305 | -31.8843
21.4809 | -35.2308 | -31.8918
21.487 | -35.2313 | -31.8978
21.4931 | -35.2317 | -31.904
21.4992 | -35.2321 | -31.9111
21.5053 | -35.2325 | -31.9169
21.5114 | -35.2329 | -31.924
21.5175 | -35.2333 | -31.9306
21.5236 | -35.2339 | -31.9367
21.5297 | -35.2345 | -31.9416
21.5358 | -35.2349 | -31.9475
21.5419 | -35.2353 | -31.9542
21.548 | -35.2359 | -31.9606
21.5541 | -35.2363 | -31.9672
21.5602 | -35.2368 | -31.9736
21.5664 | -35.2372 | -31.98
21.5725 | -35.2376 | -31.9873
21.5786 | -35.2381 | -31.9937
21.5847 | -35.2386 | -31.9997
21.5908 | -35.239 | -32.0064
21.5969 | -35.2395 | -32.0118
21.603 | -35.2399 | -32.0187
21.6091 | -35.2404 | -32.0247
21.6152 | -35.2408 | -32.031
21.6213 | -35.2413 | -32.0373
21.6274 | -35.2418 | -32.0446
21.6335 | -35.2423 | -32.051
21.6396 | -35.2426 | -32.0572
21.6457 | -35.2431 | -32.063
21.6518 | -35.2435 | -32.0682
21.6579 | -35.2438 | -32.0748
21.664 | -35.2442 | -32.0811
21.6701 | -35.2449 | -32.0871
21.6762 | -35.2454 | -32.0944
21.6823 | -35.2459 | -32.1004
21.6884 | -35.2463 | -32.1072
21.6945 | -35.2467 | -32.1136
21.7006 | -35.247 | -32.1184
21.7068 | -35.2474 | -32.1239
21.7129 | -35.248 | -32.1297
21.719 | -35.2484 | -32.1363
21.7251 | -35.2489 | -32.1426
21.7312 | -35.2494 | -32.1484
21.7373 | -35.2499 | -32.1549
21.7434 | -35.2505 | -32.1612
21.7495 | -35.2509 | -32.1681
21.7556 | -35.2515 | -32.1751
21.7617 | -35.2519 | -32.1814
21.7678 | -35.2523 | -32.1871
21.7739 | -35.2527 | -32.1943
21.78 | -35.2532 | -32.2009
21.7861 | -35.2538 | -32.2075
21.7922 | -35.2543 | -32.2142
21.7983 | -35.2547 | -32.2205
21.8044 | -35.2551 | -32.227
21.8105 | -35.2556 | -32.2313
21.8166 | -35.256 | -32.2384
21.8227 | -35.2565 | -32.2456
21.8288 | -35.2571 | -32.2514
21.8349 | -35.2576 | -32.2578
21.841 | -35.258 | -32.2642
21.8471 | -35.2585 | -32.2706
21.8533 | -35.2591 | -32.2762
21.8594 | -35.2597 | -32.2826
21.8655 | -35.2601 | -32.2877
21.8716 | -35.2606 | -32.2935
21.8777 | -35.2611 | -32.2994
21.8838 | -35.2614 | -32.3056
21.8899 | -35.2619 | -32.3118
21.896 | -35.2624 | -32.3185
21.9021 | -35.263 | -32.3256
21.9082 | -35.2634 | -32.3316
21.9143 | -35.2637 | -32.3378
21.9204 | -35.2642 | -32.3429
21.9265 | -35.2645 | -32.3491
21.9326 | -35.2648 | -32.3549
21.9387 | -35.2654 | -32.3611
21.9448 | -35.2657 | -32.3675
21.9509 | -35.266 | -32.3736
21.957 | -35.2666 | -32.3795
21.9631 | -35.267 | -32.386
21.9692 | -35.2675 | -32.3922
21.9753 | -35.268 | -32.3988
21.9814 | -35.2684 | -32.4052
21.9875 | -35.269 | -32.4104
21.9937 | -35.2695 | -32.4169
21.9998 | -35.27 | -32.4229
22.0059 | -35.2704 | -32.4293
22.012 | -35.2709 | -32.4359
22.0181 | -35.2712 | -32.442
22.0242 | -35.2716 | -32.4484
22.0303 | -35.272 | -32.4548
22.0364 | -35.2725 | -32.4604
22.0425 | -35.273 | -32.4659
22.0486 | -35.2734 | -32.4722
22.0547 | -35.2738 | -32.4787
22.0608 | -35.2741 | -32.4856
22.0669 | -35.2745 | -32.4914
22.073 | -35.275 | -32.497
22.0791 | -35.2754 | -32.5026
22.0852 | -35.276 | -32.5087
22.0913 | -35.2762 | -32.5153
22.0974 | -35.2767 | -32.5226
22.1035 | -35.2772 | -32.5286
22.1096 | -35.2777 | -32.5358
22.1157 | -35.2781 | -32.542
22.1218 | -35.2786 | -32.548
22.1279 | -35.2789 | -32.5534
22.134 | -35.2793 | -32.5596
22.1402 | -35.2797 | -32.5651
22.1463 | -35.28 | -32.572
22.1524 | -35.2804 | -32.5791
22.1585 | -35.2808 | -32.5852
22.1646 | -35.2811 | -32.5918
22.1707 | -35.2815 | -32.5984
22.1768 | -35.2821 | -32.6051
22.1829 | -35.2825 | -32.6109
22.189 | -35.2831 | -32.6177
22.1951 | -35.2837 | -32.6238
22.2012 | -35.2842 | -32.6305
22.2073 | -35.2845 | -32.6365
22.2134 | -35.2848 | -32.6428
22.2195 | -35.2852 | -32.6482
22.2256 | -35.2858 | -32.6547
22.2317 | -35.2864 | -32.6609
22.2378 | -35.2868 | -32.6665
22.2439 | -35.2874 | -32.674
22.25 | -35.2878 | -32.6809
22.2561 | -35.2884 | -32.6872
22.2622 | -35.2889 | -32.6938
22.2683 | -35.2895 | -32.6993
22.2744 | -35.2899 | -32.7058
22.2806 | -35.2903 | -32.712
22.2867 | -35.2907 | -32.7191
22.2928 | -35.2911 | -32.725
22.2989 | -35.2916 | -32.7316
22.305 | -35.2921 | -32.7396
22.3111 | -35.2925 | -32.7458
22.3172 | -35.293 | -32.7514
22.3233 | -35.2936 | -32.7577
22.3294 | -35.2941 | -32.7642
22.3355 | -35.2946 | -32.77
22.3416 | -35.2952 | -32.7761
22.3477 | -35.2957 | -32.7827
22.3538 | -35.2962 | -32.7886
22.3599 | -35.2966 | -32.7951
22.366 | -35.297 | -32.8014
22.3721 | -35.2975 | -32.8086
22.3782 | -35.2979 | -32.8154
22.3843 | -35.2982 | -32.8218
22.3904 | -35.2987 | -32.828
22.3965 | -35.2992 | -32.834
22.4026 | -35.2997 | -32.8404
22.4087 | -35.3003 | -32.8474
22.4148 | -35.3006 | -32.8541
22.4209 | -35.301 | -32.861
22.4271 | -35.3013 | -32.868
22.4332 | -35.3018 | -32.8748
22.4393 | -35.3023 | -32.8811
22.4454 | -35.3028 | -32.8876
22.4515 | -35.3032 | -32.8947
22.4576 | -35.3037 | -32.901
22.4637 | -35.3042 | -32.9068
22.4698 | -35.3047 | -32.9134
22.4759 | -35.3051 | -32.9196
22.482 | -35.3055 | -32.9259
22.4881 | -35.306 | -32.9323
22.4942 | -35.3065 | -32.9381
22.5003 | -35.307 | -32.9446
22.5064 | -35.3074 | -32.9518
22.5125 | -35.3078 | -32.9571
22.5186 | -35.3082 | -32.9647
22.5247 | -35.3087 | -32.9713
22.5308 | -35.3092 | -32.9773
22.5369 | -35.3097 | -32.9833
22.543 | -35.3102 | -32.9899
22.5491 | -35.3106 | -32.9967
22.5552 | -35.3109 | -33.0029
22.5613 | -35.3114 | -33.0094
22.5675 | -35.3118 | -33.0152
22.5736 | -35.3124 | -33.0215
22.5797 | -35.3127 | -33.0288
22.5858 | -35.3132 | -33.0354
22.5919 | -35.3137 | -33.0412
22.598 | -35.3141 | -33.0471
22.6041 | -35.3146 | -33.0528
22.6102 | -35.315 | -33.0597
22.6163 | -35.3155 | -33.0648
22.6224 | -35.316 | -33.0707
22.6285 | -35.3165 | -33.0771
22.6346 | -35.3169 | -33.0837
22.6407 | -35.3174 | -33.0897
22.6468 | -35.3179 | -33.0956
22.6529 | -35.3184 | -33.1015
22.659 | -35.3187 | -33.1073
22.6651 | -35.3191 | -33.1136
22.6712 | -35.3196 | -33.1193
22.6773 | -35.3201 | -33.125
22.6834 | -35.3206 | -33.1314
22.6895 | -35.3211 | -33.1374
22.6956 | -35.3216 | -33.1445
22.7017 | -35.3221 | -33.151
22.7079 | -35.3227 | -33.1573
22.714 | -35.3232 | -33.1627
22.7201 | -35.3237 | -33.1696
22.7262 | -35.3241 | -33.1762
22.7323 | -35.3246 | -33.1827
22.7384 | -35.3251 | -33.1885
22.7445 | -35.3256 | -33.1955
22.7506 | -35.326 | -33.2016
22.7567 | -35.3266 | -33.2074
22.7628 | -35.3272 | -33.2135
22.7689 | -35.3276 | -33.2197
22.775 | -35.3281 | -33.2258
22.7811 | -35.3286 | -33.232
22.7872 | -35.3291 | -33.2378
22.7933 | -35.3295 | -33.2442
22.7994 | -35.3299 | -33.2499
22.8055 | -35.3304 | -33.2557
22.8116 | -35.3308 | -33.2621
22.8177 | -35.3314 | -33.2681
22.8238 | -35.3319 | -33.2746
22.8299 | -35.3324 | -33.2812
22.836 | -35.3329 | -33.2874
22.8421 | -35.3334 | -33.2936
22.8482 | -35.3338 | -33.3004
22.8544 | -35.3343 | -33.3065
22.8605 | -35.3347 | -33.313
22.8666 | -35.3352 | -33.3197
22.8727 | -35.3358 | -33.3257
22.8788 | -35.3362 | -33.3319
22.8849 | -35.3366 | -33.3376
22.891 | -35.337 | -33.3438
22.8971 | -35.3375 | -33.3498
22.9032 | -35.338 | -33.3556
22.9093 | -35.3385 | -33.3621
22.9154 | -35.339 | -33.3686
22.9215 | -35.3396 | -33.3748
22.9276 | -35.3401 | -33.3809
22.9337 | -35.3405 | -33.3879
22.9398 | -35.3408 | -33.3945
22.9459 | -35.3414 | -33.4002
22.952 | -35.3419 | -33.407
22.9581 | -35.3423 | -33.413
22.9642 | -35.3428 | -33.4187
22.9703 | -35.3432 | -33.4254
22.9764 | -35.3436 | -33.4317
22.9825 | -35.3441 | -33.4377
22.9886 | -35.3447 | -33.4444
22.9948 | -35.3451 | -33.4508
23.0009 | -35.3456 | -33.4573
23.007 | -35.3463 | -33.4638
23.0131 | -35.3468 | -33.4705
23.0192 | -35.3472 | -33.4761
23.0253 | -35.3477 | -33.4824
23.0314 | -35.3482 | -33.4883
23.0375 | -35.3486 | -33.4945
23.0436 | -35.349 | -33.5012
23.0497 | -35.3494 | -33.507
23.0558 | -35.3498 | -33.5133
23.0619 | -35.3502 | -33.5191
23.068 | -35.3507 | -33.5249
23.0741 | -35.3512 | -33.5309
23.0802 | -35.3517 | -33.5372
23.0863 | -35.3522 | -33.5435
23.0924 | -35.3527 | -33.5494
23.0985 | -35.3532 | -33.5554
23.1046 | -35.3537 | -33.5613
23.1107 | -35.3541 | -33.5677
23.1168 | -35.3546 | -33.5739
23.1229 | -35.3551 | -33.5801
23.129 | -35.3555 | -33.586
23.1351 | -35.3559 | -33.5923
23.1413 | -35.3564 | -33.598
23.1474 | -35.3568 | -33.6045
23.1535 | -35.3573 | -33.6107
23.1596 | -35.3578 | -33.6169
23.1657 | -35.3582 | -33.6229
23.1718 | -35.3587 | -33.6281
23.1779 | -35.3591 | -33.6343
23.184 | -35.3596 | -33.6401
23.1901 | -35.3601 | -33.646
23.1962 | -35.3605 | -33.6523
23.2023 | -35.361 | -33.6581
23.2084 | -35.3615 | -33.6646
23.2145 | -35.362 | -33.6711
23.2206 | -35.3625 | -33.6781
23.2267 | -35.3629 | -33.6839
23.2328 | -35.3634 | -33.6902
23.2389 | -35.3639 | -33.6963
23.245 | -35.3644 | -33.7022
23.2511 | -35.3648 | -33.7083
23.2572 | -35.3652 | -33.7142
23.2633 | -35.3658 | -33.7215
23.2694 | -35.3663 | -33.7286
23.2755 | -35.3668 | -33.735
23.2817 | -35.3673 | -33.7414
23.2878 | -35.3678 | -33.747
23.2939 | -35.3684 | -33.7531
23.3 | -35.3688 | -33.7588
23.3061 | -35.3693 | -33.7652
23.3122 | -35.3698 | -33.771
23.3183 | -35.3703 | -33.7773
23.3244 | -35.3707 | -33.7844
23.3305 | -35.3712 | -33.7904
23.3366 | -35.3717 | -33.7966
23.3427 | -35.3722 | -33.8029
23.3488 | -35.3727 | -33.809
23.3549 | -35.3732 | -33.8149
23.361 | -35.3739 | -33.8207
23.3671 | -35.3743 | -33.8267
23.3732 | -35.3749 | -33.8324
23.3793 | -35.3754 | -33.8383
23.3854 | -35.3759 | -33.8445
23.3915 | -35.3765 | -33.8509
23.3976 | -35.377 | -33.8569
23.4037 | -35.3775 | -33.8631
23.4098 | -35.3781 | -33.8698
23.4159 | -35.3785 | -33.8759
23.422 | -35.3789 | -33.8818
23.4282 | -35.3794 | -33.8873
23.4343 | -35.3798 | -33.894
23.4404 | -35.3803 | -33.9005
23.4465 | -35.3808 | -33.9072
23.4526 | -35.3813 | -33.9139
23.4587 | -35.3818 | -33.9201
23.4648 | -35.3824 | -33.9264
23.4709 | -35.3828 | -33.9327
23.477 | -35.3834 | -33.9387
23.4831 | -35.3839 | -33.9452
23.4892 | -35.3843 | -33.9515
23.4953 | -35.3847 | -33.9575
23.5014 | -35.3853 | -33.9638
23.5075 | -35.3857 | -33.9702
23.5136 | -35.3862 | -33.9764
23.5197 | -35.3867 | -33.9827
23.5258 | -35.3872 | -33.9887
23.5319 | -35.3877 | -33.9945
23.538 | -35.3882 | -34.0003
23.5441 | -35.3887 | -34.0061
23.5502 | -35.3892 | -34.0127
23.5563 | -35.3897 | -34.0191
23.5624 | -35.3901 | -34.0263
23.5686 | -35.3906 | -34.0331
23.5747 | -35.3911 | -34.0393
23.5808 | -35.3916 | -34.0461
23.5869 | -35.3921 | -34.0521
23.593 | -35.3926 | -34.0586
23.5991 | -35.393 | -34.0645
23.6052 | -35.3935 | -34.0706
23.6113 | -35.394 | -34.0768
23.6174 | -35.3945 | -34.0828
23.6235 | -35.3949 | -34.0886
23.6296 | -35.3953 | -34.0951
23.6357 | -35.3959 | -34.1014
23.6418 | -35.3965 | -34.1071
23.6479 | -35.3969 | -34.1131
23.654 | -35.3974 | -34.1192
23.6601 | -35.3979 | -34.1252
23.6662 | -35.3983 | -34.1318
23.6723 | -35.3989 | -34.138
23.6784 | -35.3994 | -34.144
23.6845 | -35.3998 | -34.1499
23.6906 | -35.4003 | -34.1561
23.6967 | -35.4008 | -34.1615
23.7028 | -35.4012 | -34.1676
23.7089 | -35.4016 | -34.174
23.7151 | -35.4021 | -34.1801
23.7212 | -35.4025 | -34.1868
23.7273 | -35.4029 | -34.1929
23.7334 | -35.4034 | -34.1997
23.7395 | -35.404 | -34.2056
23.7456 | -35.4044 | -34.2124
23.7517 | -35.4049 | -34.2184
23.7578 | -35.4053 | -34.2247
23.7639 | -35.4058 | -34.2309
23.77 | -35.4062 | -34.2367
23.7761 | -35.4066 | -34.2433
23.7822 | -35.4071 | -34.2491
23.7883 | -35.4076 | -34.2553
23.7944 | -35.408 | -34.2621
23.8005 | -35.4085 | -34.269
23.8066 | -35.4091 | -34.2749
23.8127 | -35.4096 | -34.2813
23.8188 | -35.4101 | -34.2873
23.8249 | -35.4107 | -34.2933
23.831 | -35.4111 | -34.2992
23.8371 | -35.4116 | -34.3054
23.8432 | -35.412 | -34.3114
23.8493 | -35.4124 | -34.3173
23.8555 | -35.4128 | -34.3238
23.8616 | -35.4133 | -34.33
23.8677 | -35.4136 | -34.3363
23.8738 | -35.4141 | -34.3422
23.8799 | -35.4147 | -34.3487
23.886 | -35.4151 | -34.3544
23.8921 | -35.4156 | -34.3601
23.8982 | -35.416 | -34.3656
23.9043 | -35.4165 | -34.3718
23.9104 | -35.4169 | -34.378
23.9165 | -35.4175 | -34.3839
23.9226 | -35.4179 | -34.3898
23.9287 | -35.4184 | -34.3964
23.9348 | -35.4188 | -34.4026
23.9409 | -35.4193 | -34.4092
23.947 | -35.4198 | -34.4157
23.9531 | -35.4202 | -34.4221
23.9592 | -35.4207 | -34.4287
23.9653 | -35.4212 | -34.4354
23.9714 | -35.4216 | -34.4411
23.9775 | -35.422 | -34.4473
23.9836 | -35.4225 | -34.4532
23.9897 | -35.4231 | -34.4594
23.9958 | -35.4235 | -34.4655
24.002 | -35.4239 | -34.4723
24.0081 | -35.4244 | -34.4782
24.0142 | -35.4249 | -34.4845
24.0203 | -35.4254 | -34.4909
24.0264 | -35.4258 | -34.4973
24.0325 | -35.4263 | -34.5038
24.0386 | -35.4266 | -34.51
24.0447 | -35.4271 | -34.5161
24.0508 | -35.4276 | -34.5218
24.0569 | -35.4281 | -34.5281
24.063 | -35.4285 | -34.5342
24.0691 | -35.429 | -34.5407
24.0752 | -35.4294 | -34.5472
24.0813 | -35.4298 | -34.5535
24.0874 | -35.4302 | -34.5595
24.0935 | -35.4306 | -34.5651
24.0996 | -35.4311 | -34.5721
24.1057 | -35.4315 | -34.5778
24.1118 | -35.432 | -34.584
24.1179 | -35.4325 | -34.5904
24.124 | -35.433 | -34.5955
24.1301 | -35.4335 | -34.6014
24.1362 | -35.434 | -34.608
24.1424 | -35.4345 | -34.6146
24.1485 | -35.4349 | -34.6203
24.1546 | -35.4354 | -34.6262
24.1607 | -35.4359 | -34.6327
24.1668 | -35.4363 | -34.6381
24.1729 | -35.4367 | -34.6444
24.179 | -35.4371 | -34.6504
24.1851 | -35.4375 | -34.6564
24.1912 | -35.4379 | -34.6628
24.1973 | -35.4384 | -34.6698
24.2034 | -35.439 | -34.6762
24.2095 | -35.4395 | -34.6823
24.2156 | -35.44 | -34.6881
24.2217 | -35.4404 | -34.6942
24.2278 | -35.4409 | -34.7004
24.2339 | -35.4414 | -34.7069
24.24 | -35.4419 | -34.713
24.2461 | -35.4424 | -34.7191
24.2522 | -35.4429 | -34.7256
24.2583 | -35.4433 | -34.7321
24.2644 | -35.4438 | -34.7376
24.2705 | -35.4442 | -34.7441
24.2766 | -35.4447 | -34.7506
24.2827 | -35.4451 | -34.7569
24.2889 | -35.4456 | -34.7631
24.295 | -35.446 | -34.7693
24.3011 | -35.4465 | -34.7753
24.3072 | -35.4469 | -34.7817
24.3133 | -35.4474 | -34.7878
24.3194 | -35.4478 | -34.7941
24.3255 | -35.4482 | -34.8002
24.3316 | -35.4488 | -34.8065
24.3377 | -35.4492 | -34.8122
24.3438 | -35.4496 | -34.8186
24.3499 | -35.4501 | -34.8245
24.356 | -35.4505 | -34.8306
24.3621 | -35.4509 | -34.8367
24.3682 | -35.4514 | -34.8428
24.3743 | -35.4518 | -34.8494
24.3804 | -35.4523 | -34.8553
24.3865 | -35.4527 | -34.8615
24.3926 | -35.4532 | -34.8675
24.3987 | -35.4537 | -34.8738
24.4048 | -35.4542 | -34.8803
24.4109 | -35.4546 | -34.8865
24.417 | -35.455 | -34.8929
24.4231 | -35.4555 | -34.8994
24.4293 | -35.4558 | -34.9057
24.4354 | -35.4563 | -34.9117
24.4415 | -35.4568 | -34.9178
24.4476 | -35.4572 | -34.9243
24.4537 | -35.4577 | -34.9306
24.4598 | -35.4582 | -34.9369
24.4659 | -35.4586 | -34.9429
24.472 | -35.459 | -34.9489
24.4781 | -35.4595 | -34.9551
24.4842 | -35.46 | -34.9614
24.4903 | -35.4605 | -34.9672
24.4964 | -35.4609 | -34.9735
24.5025 | -35.4614 | -34.9793
24.5086 | -35.4618 | -34.9858
24.5147 | -35.4623 | -34.9919
24.5208 | -35.4627 | -34.9977
24.5269 | -35.4632 | -35.0039
24.533 | -35.4637 | -35.0102
24.5391 | -35.4641 | -35.0165
24.5452 | -35.4645 | -35.0223
24.5513 | -35.4649 | -35.0287
24.5574 | -35.4654 | -35.0349
24.5635 | -35.4659 | -35.0413
24.5696 | -35.4663 | -35.0468
24.5758 | -35.4667 | -35.0527
24.5819 | -35.4672 | -35.0589
24.588 | -35.4676 | -35.0647
24.5941 | -35.4681 | -35.0707
24.6002 | -35.4686 | -35.0767
24.6063 | -35.4692 | -35.0827
24.6124 | -35.4696 | -35.0887
24.6185 | -35.4701 | -35.095
24.6246 | -35.4705 | -35.101
24.6307 | -35.471 | -35.1068
24.6368 | -35.4714 | -35.1132
24.6429 | -35.4719 | -35.1194
24.649 | -35.4724 | -35.1258
24.6551 | -35.4729 | -35.1318
24.6612 | -35.4733 | -35.1383
24.6673 | -35.4738 | -35.1448
24.6734 | -35.4743 | -35.1509
24.6795 | -35.4747 | -35.1571
24.6856 | -35.4752 | -35.1634
24.6917 | -35.4756 | -35.1699
24.6978 | -35.4761 | -35.1761
24.7039 | -35.4765 | -35.1822
24.71 | -35.477 | -35.1886
24.7162 | -35.4774 | -35.1947
24.7223 | -35.478 | -35.201
24.7284 | -35.4784 | -35.2068
24.7345 | -35.4789 | -35.2127
24.7406 | -35.4793 | -35.2189
24.7467 | -35.4798 | -35.2254
24.7528 | -35.4802 | -35.2315
24.7589 | -35.4806 | -35.2372
24.765 | -35.481 | -35.2437
24.7711 | -35.4815 | -35.2498
24.7772 | -35.482 | -35.2558
24.7833 | -35.4824 | -35.262
24.7894 | -35.4828 | -35.2684
24.7955 | -35.4832 | -35.2748
24.8016 | -35.4836 | -35.2807
24.8077 | -35.484 | -35.2868
24.8138 | -35.4844 | -35.2936
24.8199 | -35.4849 | -35.2997
24.826 | -35.4854 | -35.3059
24.8321 | -35.4858 | -35.3119
24.8382 | -35.4862 | -35.3182
24.8443 | -35.4867 | -35.3243
24.8504 | -35.4872 | -35.3305
24.8565 | -35.4876 | -35.3364
24.8627 | -35.4881 | -35.3423
24.8688 | -35.4884 | -35.348
24.8749 | -35.4889 | -35.3538
24.881 | -35.4893 | -35.3598
24.8871 | -35.4898 | -35.3661
24.8932 | -35.4903 | -35.3724
24.8993 | -35.4906 | -35.3782
24.9054 | -35.4911 | -35.3843
24.9115 | -35.4915 | -35.3905
24.9176 | -35.4919 | -35.3963
24.9237 | -35.4923 | -35.4023
24.9298 | -35.4928 | -35.4086
24.9359 | -35.4932 | -35.415
24.942 | -35.4935 | -35.4209
24.9481 | -35.4939 | -35.4271
24.9542 | -35.4943 | -35.4333
24.9603 | -35.4947 | -35.4394
24.9664 | -35.4952 | -35.4457
24.9725 | -35.4956 | -35.4522
24.9786 | -35.496 | -35.4579
24.9847 | -35.4963 | -35.4637
24.9908 | -35.4968 | -35.4704
24.9969 | -35.4972 | -35.4764
25.0031 | -35.4977 | -35.4827
25.0092 | -35.4981 | -35.4886
25.0153 | -35.4985 | -35.4947
25.0214 | -35.4989 | -35.5007
25.0275 | -35.4993 | -35.5063
25.0336 | -35.4998 | -35.5121
25.0397 | -35.5002 | -35.5185
25.0458 | -35.5006 | -35.5246
25.0519 | -35.5011 | -35.5305
25.058 | -35.5015 | -35.5368
25.0641 | -35.5019 | -35.5431
25.0702 | -35.5023 | -35.5494
25.0763 | -35.5028 | -35.5555
25.0824 | -35.5032 | -35.5613
25.0885 | -35.5036 | -35.5675
25.0946 | -35.5041 | -35.5736
25.1007 | -35.5045 | -35.5798
25.1068 | -35.5049 | -35.5857
25.1129 | -35.5054 | -35.592
25.119 | -35.5058 | -35.5979
25.1251 | -35.5062 | -35.6039
25.1312 | -35.5067 | -35.6098
25.1373 | -35.5071 | -35.616
25.1435 | -35.5075 | -35.6223
25.1496 | -35.5078 | -35.6284
25.1557 | -35.5082 | -35.6346
25.1618 | -35.5087 | -35.641
25.1679 | -35.5091 | -35.6472
25.174 | -35.5095 | -35.6527
25.1801 | -35.51 | -35.659
25.1862 | -35.5104 | -35.6647
25.1923 | -35.5108 | -35.6707
25.1984 | -35.5113 | -35.6771
25.2045 | -35.5118 | -35.6831
25.2106 | -35.5122 | -35.6896
25.2167 | -35.5126 | -35.6952
25.2228 | -35.513 | -35.7013
25.2289 | -35.5134 | -35.7073
25.235 | -35.5138 | -35.7136
25.2411 | -35.5142 | -35.7195
25.2472 | -35.5146 | -35.7257
25.2533 | -35.515 | -35.7316
25.2594 | -35.5154 | -35.7379
25.2655 | -35.5158 | -35.744
25.2716 | -35.5162 | -35.7505
25.2777 | -35.5166 | -35.7568
25.2838 | -35.517 | -35.763
25.29 | -35.5174 | -35.7692
25.2961 | -35.5178 | -35.7753
25.3022 | -35.5182 | -35.7812
25.3083 | -35.5186 | -35.7871
25.3144 | -35.519 | -35.7929
25.3205 | -35.5195 | -35.799
25.3266 | -35.5198 | -35.805
25.3327 | -35.5203 | -35.8112
25.3388 | -35.5206 | -35.8172
25.3449 | -35.521 | -35.8233
25.351 | -35.5214 | -35.8293
25.3571 | -35.5218 | -35.8355
25.3632 | -35.5221 | -35.8418
25.3693 | -35.5224 | -35.8478
25.3754 | -35.5228 | -35.8536
25.3815 | -35.5233 | -35.8598
25.3876 | -35.5236 | -35.866
25.3937 | -35.524 | -35.8718
25.3998 | -35.5244 | -35.8773
25.4059 | -35.5248 | -35.8838
25.412 | -35.5252 | -35.8897
25.4181 | -35.5256 | -35.8959
25.4242 | -35.5261 | -35.902
25.4304 | -35.5265 | -35.908
25.4365 | -35.5269 | -35.9138
25.4426 | -35.5272 | -35.9194
25.4487 | -35.5277 | -35.9255
25.4548 | -35.528 | -35.9316
25.4609 | -35.5285 | -35.9374
25.467 | -35.5288 | -35.9429
25.4731 | -35.5292 | -35.9488
25.4792 | -35.5296 | -35.955
25.4853 | -35.53 | -35.9612
25.4914 | -35.5304 | -35.967
25.4975 | -35.5308 | -35.9729
25.5036 | -35.5311 | -35.9787
25.5097 | -35.5316 | -35.9848
25.5158 | -35.532 | -35.9907
25.5219 | -35.5324 | -35.9969
25.528 | -35.5328 | -36.0029
25.5341 | -35.5332 | -36.009
25.5402 | -35.5335 | -36.0149
25.5463 | -35.5339 | -36.0205
25.5524 | -35.5343 | -36.0263
25.5585 | -35.5347 | -36.0323
25.5646 | -35.5351 | -36.0381
25.5707 | -35.5355 | -36.0441
25.5769 | -35.5358 | -36.0498
25.583 | -35.5363 | -36.0559
25.5891 | -35.5367 | -36.0618
25.5952 | -35.5371 | -36.0678
25.6013 | -35.5375 | -36.0741
25.6074 | -35.5378 | -36.0798
25.6135 | -35.5382 | -36.0856
25.6196 | -35.5385 | -36.092
25.6257 | -35.5389 | -36.098
25.6318 | -35.5394 | -36.1041
25.6379 | -35.5397 | -36.1101
25.644 | -35.54 | -36.1158
25.6501 | -35.5404 | -36.1217
25.6562 | -35.5408 | -36.128
25.6623 | -35.5412 | -36.134
25.6684 | -35.5415 | -36.1402
25.6745 | -35.5419 | -36.1465
25.6806 | -35.5423 | -36.1522
25.6867 | -35.5427 | -36.1583
25.6928 | -35.5431 | -36.1643
25.6989 | -35.5434 | -36.1702
25.705 | -35.5438 | -36.1765
25.7111 | -35.5442 | -36.1823
25.7173 | -35.5445 | -36.1877
25.7234 | -35.5449 | -36.1937
25.7295 | -35.5453 | -36.1999
25.7356 | -35.5456 | -36.206
25.7417 | -35.546 | -36.2118
25.7478 | -35.5463 | -36.2169
25.7539 | -35.5467 | -36.2224
25.76 | -35.5471 | -36.2282
25.7661 | -35.5475 | -36.2338
25.7722 | -35.5479 | -36.2399
25.7783 | -35.5483 | -36.2457
25.7844 | -35.5486 | -36.251
25.7905 | -35.549 | -36.2569
25.7966 | -35.5494 | -36.2627
25.8027 | -35.5497 | -36.2684
25.8088 | -35.5501 | -36.2743
25.8149 | -35.5504 | -36.2801
25.821 | -35.5508 | -36.2858
25.8271 | -35.5512 | -36.2917
25.8332 | -35.5516 | -36.2974
25.8393 | -35.5519 | -36.3032
25.8454 | -35.5523 | -36.3085
25.8515 | -35.5527 | -36.3143
25.8576 | -35.553 | -36.3197
25.8638 | -35.5535 | -36.3259
25.8699 | -35.5538 | -36.3317
25.876 | -35.5541 | -36.3377
25.8821 | -35.5544 | -36.3435
25.8882 | -35.5547 | -36.3493
25.8943 | -35.5552 | -36.3554
25.9004 | -35.5555 | -36.3614
25.9065 | -35.5559 | -36.3674
25.9126 | -35.5562 | -36.3733
25.9187 | -35.5565 | -36.3792
25.9248 | -35.5569 | -36.3845
25.9309 | -35.5573 | -36.3905
25.937 | -35.5577 | -36.3964
25.9431 | -35.558 | -36.4021
25.9492 | -35.5584 | -36.4081
25.9553 | -35.5588 | -36.414
25.9614 | -35.5591 | -36.4197
25.9675 | -35.5596 | -36.4256
25.9736 | -35.56 | -36.4317
25.9797 | -35.5603 | -36.4374
25.9858 | -35.5607 | -36.443
25.9919 | -35.5611 | -36.4486
25.998 | -35.5614 | -36.4542
26.0042 | -35.5617 | -36.4601
26.0103 | -35.5621 | -36.466
26.0164 | -35.5624 | -36.4719
26.0225 | -35.5628 | -36.4777
26.0286 | -35.5631 | -36.4836
26.0347 | -35.5634 | -36.4891
26.0408 | -35.5637 | -36.4953
26.0469 | -35.5641 | -36.5012
26.053 | -35.5644 | -36.507
26.0591 | -35.5648 | -36.5127
26.0652 | -35.5651 | -36.518
26.0713 | -35.5654 | -36.5237
26.0774 | -35.5658 | -36.5295
26.0835 | -35.5661 | -36.5354
26.0896 | -35.5666 | -36.5411
26.0957 | -35.5669 | -36.5467
26.1018 | -35.5673 | -36.5521
26.1079 | -35.5676 | -36.5585
26.114 | -35.5679 | -36.5644
26.1201 | -35.5683 | -36.5703
26.1262 | -35.5686 | -36.5762
26.1323 | -35.569 | -36.5819
26.1384 | -35.5693 | -36.5875
26.1445 | -35.5696 | -36.5934
26.1507 | -35.5699 | -36.5992
26.1568 | -35.5703 | -36.6051
26.1629 | -35.5706 | -36.611
26.169 | -35.571 | -36.6169
26.1751 | -35.5713 | -36.6226
26.1812 | -35.5717 | -36.6286
26.1873 | -35.5721 | -36.6346
26.1934 | -35.5724 | -36.6402
26.1995 | -35.5727 | -36.6462
26.2056 | -35.573 | -36.6515
26.2117 | -35.5733 | -36.6573
26.2178 | -35.5737 | -36.6629
26.2239 | -35.574 | -36.6687
26.23 | -35.5744 | -36.6746
26.2361 | -35.5747 | -36.6801
26.2422 | -35.575 | -36.6854
26.2483 | -35.5754 | -36.691
26.2544 | -35.5758 | -36.6965
26.2605 | -35.5761 | -36.7023
26.2666 | -35.5765 | -36.708
26.2727 | -35.5768 | -36.7141
26.2788 | -35.5771 | -36.7197
26.2849 | -35.5776 | -36.7253
26.2911 | -35.578 | -36.7309
26.2972 | -35.5784 | -36.7367
26.3033 | -35.5787 | -36.7424
26.3094 | -35.5791 | -36.748
26.3155 | -35.5794 | -36.7531
26.3216 | -35.5798 | -36.7587
26.3277 | -35.5801 | -36.7641
26.3338 | -35.5805 | -36.7697
26.3399 | -35.5809 | -36.7751
26.346 | -35.5813 | -36.7801
26.3521 | -35.5816 | -36.7861
26.3582 | -35.5819 | -36.7921
26.3643 | -35.5822 | -36.7977
26.3704 | -35.5825 | -36.8034
26.3765 | -35.5829 | -36.8088
26.3826 | -35.5832 | -36.8139
26.3887 | -35.5835 | -36.8195
26.3948 | -35.5839 | -36.8247
26.4009 | -35.5842 | -36.8303
26.407 | -35.5846 | -36.8362
26.4131 | -35.5849 | -36.8421
26.4192 | -35.5852 | -36.8475
26.4253 | -35.5855 | -36.8535
26.4314 | -35.5857 | -36.8593
26.4376 | -35.5861 | -36.865
26.4437 | -35.5864 | -36.8707
26.4498 | -35.5867 | -36.8762
26.4559 | -35.587 | -36.8811
26.462 | -35.5874 | -36.8868
26.4681 | -35.5877 | -36.8921
26.4742 | -35.588 | -36.898
26.4803 | -35.5883 | -36.9032
26.4864 | -35.5886 | -36.9086
26.4925 | -35.5889 | -36.9141
26.4986 | -35.5893 | -36.9195
26.5047 | -35.5896 | -36.9255
26.5108 | -35.59 | -36.9309
26.5169 | -35.5903 | -36.9364
26.523 | -35.5907 | -36.9417
26.5291 | -35.591 | -36.947
26.5352 | -35.5914 | -36.9525
26.5413 | -35.5917 | -36.9581
26.5474 | -35.592 | -36.964
26.5535 | -35.5924 | -36.9697
26.5596 | -35.5926 | -36.9751
26.5657 | -35.593 | -36.9808
26.5718 | -35.5933 | -36.986
26.578 | -35.5936 | -36.9916
26.5841 | -35.5939 | -36.9972
26.5902 | -35.5942 | -37.0024
26.5963 | -35.5945 | -37.0074
26.6024 | -35.5948 | -37.0133
26.6085 | -35.5952 | -37.0193
26.6146 | -35.5955 | -37.0247
26.6207 | -35.5958 | -37.0307
26.6268 | -35.5961 | -37.0361
26.6329 | -35.5964 | -37.0415
26.639 | -35.5968 | -37.0473
26.6451 | -35.5971 | -37.0526
26.6512 | -35.5974 | -37.058
26.6573 | -35.5978 | -37.0636
26.6634 | -35.5981 | -37.0691
26.6695 | -35.5983 | -37.0747
26.6756 | -35.5987 | -37.0804
26.6817 | -35.599 | -37.0861
26.6878 | -35.5993 | -37.0918
26.6939 | -35.5997 | -37.0977
26.7 | -35.5999 | -37.103
26.7061 | -35.6002 | -37.1083
26.7122 | -35.6005 | -37.1141
26.7183 | -35.6009 | -37.1191
26.7245 | -35.6013 | -37.1247
26.7306 | -35.6015 | -37.1302
26.7367 | -35.6018 | -37.1351
26.7428 | -35.6021 | -37.141
26.7489 | -35.6024 | -37.1464
26.755 | -35.6027 | -37.1521
26.7611 | -35.6031 | -37.1576
26.7672 | -35.6034 | -37.1633
26.7733 | -35.6037 | -37.1685
26.7794 | -35.604 | -37.1742
26.7855 | -35.6043 | -37.1801
26.7916 | -35.6046 | -37.1859
26.7977 | -35.605 | -37.1912
26.8038 | -35.6053 | -37.1968
26.8099 | -35.6056 | -37.2024
26.816 | -35.6058 | -37.2081
26.8221 | -35.6062 | -37.2136
26.8282 | -35.6064 | -37.2196
26.8343 | -35.6068 | -37.2252
26.8404 | -35.6072 | -37.2307
26.8465 | -35.6075 | -37.2362
26.8526 | -35.6079 | -37.2418
26.8587 | -35.6081 | -37.2473
26.8649 | -35.6086 | -37.2525
26.871 | -35.6088 | -37.258
26.8771 | -35.6091 | -37.2631
26.8832 | -35.6095 | -37.2682
26.8893 | -35.6098 | -37.2736
26.8954 | -35.6101 | -37.2794
26.9015 | -35.6104 | -37.2849
26.9076 | -35.6107 | -37.2905
26.9137 | -35.611 | -37.2959
26.9198 | -35.6113 | -37.3017
26.9259 | -35.6116 | -37.3075
26.932 | -35.6119 | -37.3127
26.9381 | -35.6122 | -37.3183
26.9442 | -35.6125 | -37.3233
26.9503 | -35.6129 | -37.3289
26.9564 | -35.6132 | -37.3342
26.9625 | -35.6135 | -37.3398
26.9686 | -35.6138 | -37.3451
26.9747 | -35.6141 | -37.3502
26.9808 | -35.6144 | -37.3555
26.9869 | -35.6147 | -37.3609
26.993 | -35.6149 | -37.3666
26.9991 | -35.6152 | -37.3722
27.0052 | -35.6155 | -37.3775
27.0114 | -35.6158 | -37.3831
27.0175 | -35.6161 | -37.3882
27.0236 | -35.6164 | -37.394
27.0297 | -35.6167 | -37.3992
27.0358 | -35.617 | -37.4048
27.0419 | -35.6173 | -37.4104
27.048 | -35.6176 | -37.4158
27.0541 | -35.618 | -37.421
27.0602 | -35.6183 | -37.4262
27.0663 | -35.6187 | -37.4321
27.0724 | -35.619 | -37.4376
27.0785 | -35.6193 | -37.4431
27.0846 | -35.6197 | -37.4482
27.0907 | -35.62 | -37.4538
27.0968 | -35.6203 | -37.459
27.1029 | -35.6206 | -37.4645
27.109 | -35.6209 | -37.4698
27.1151 | -35.6213 | -37.4752
27.1212 | -35.6215 | -37.4806
27.1273 | -35.6218 | -37.4858
27.1334 | -35.6222 | -37.4917
27.1395 | -35.6225 | -37.4972
27.1456 | -35.6228 | -37.5032
27.1518 | -35.6231 | -37.509
27.1579 | -35.6234 | -37.5141
27.164 | -35.6238 | -37.5198
27.1701 | -35.6241 | -37.5248
27.1762 | -35.6245 | -37.5302
27.1823 | -35.6248 | -37.5358
27.1884 | -35.6251 | -37.5412
27.1945 | -35.6253 | -37.5466
27.2006 | -35.6257 | -37.5515
27.2067 | -35.6259 | -37.5567
27.2128 | -35.6263 | -37.5622
27.2189 | -35.6266 | -37.5674
27.225 | -35.6268 | -37.5727
27.2311 | -35.6271 | -37.5779
27.2372 | -35.6274 | -37.5831
27.2433 | -35.6277 | -37.5882
27.2494 | -35.6281 | -37.5936
27.2555 | -35.6284 | -37.599
27.2616 | -35.6287 | -37.6042
27.2677 | -35.629 | -37.6091
27.2738 | -35.6292 | -37.614
27.2799 | -35.6295 | -37.619
27.286 | -35.6298 | -37.6243
27.2921 | -35.6301 | -37.6294
27.2983 | -35.6304 | -37.6347
27.3044 | -35.6306 | -37.6401
27.3105 | -35.6309 | -37.6454
27.3166 | -35.6312 | -37.6505
27.3227 | -35.6315 | -37.6557
27.3288 | -35.6318 | -37.6608
27.3349 | -35.6321 | -37.6657
27.341 | -35.6325 | -37.6711
27.3471 | -35.6327 | -37.6768
27.3532 | -35.6331 | -37.6819
27.3593 | -35.6335 | -37.6873
27.3654 | -35.6338 | -37.6926
27.3715 | -35.634 | -37.6976
27.3776 | -35.6344 | -37.7035
27.3837 | -35.6347 | -37.7084
27.3898 | -35.635 | -37.7134
27.3959 | -35.6352 | -37.719
27.402 | -35.6355 | -37.7239
27.4081 | -35.6357 | -37.7291
27.4142 | -35.636 | -37.7346
27.4203 | -35.6362 | -37.7401
27.4264 | -35.6366 | -37.7452
27.4325 | -35.6368 | -37.7504
27.4387 | -35.6371 | -37.7556
27.4448 | -35.6373 | -37.761
27.4509 | -35.6375 | -37.7662
27.457 | -35.6379 | -37.7718
27.4631 | -35.6382 | -37.777
27.4692 | -35.6384 | -37.7823
27.4753 | -35.6387 | -37.7872
27.4814 | -35.639 | -37.7924
27.4875 | -35.6392 | -37.7976
27.4936 | -35.6395 | -37.8028
27.4997 | -35.6398 | -37.8083
27.5058 | -35.64 | -37.8135
27.5119 | -35.6402 | -37.8183
27.518 | -35.6405 | -37.8237
27.5241 | -35.6408 | -37.8293
27.5302 | -35.641 | -37.8344
27.5363 | -35.6413 | -37.8395
27.5424 | -35.6415 | -37.8446
27.5485 | -35.6418 | -37.8502
27.5546 | -35.6422 | -37.8556
27.5607 | -35.6424 | -37.8603
27.5668 | -35.6427 | -37.8653
27.5729 | -35.643 | -37.8706
27.5791 | -35.6433 | -37.8758
27.5852 | -35.6436 | -37.8812
27.5913 | -35.6439 | -37.8866
27.5974 | -35.6442 | -37.8915
27.6035 | -35.6445 | -37.8971
27.6096 | -35.6447 | -37.9027
27.6157 | -35.645 | -37.9076
27.6218 | -35.6453 | -37.9136
27.6279 | -35.6456 | -37.9191
27.634 | -35.6459 | -37.9246
27.6401 | -35.646 | -37.9298
27.6462 | -35.6463 | -37.935
27.6523 | -35.6466 | -37.9399
27.6584 | -35.6468 | -37.9454
27.6645 | -35.6471 | -37.9506
27.6706 | -35.6474 | -37.9561
27.6767 | -35.6477 | -37.961
27.6828 | -35.6479 | -37.9662
27.6889 | -35.6482 | -37.9715
27.695 | -35.6484 | -37.9766
27.7011 | -35.6487 | -37.9817
27.7072 | -35.649 | -37.987
27.7133 | -35.6492 | -37.9922
27.7194 | -35.6495 | -37.9972
27.7256 | -35.6498 | -38.0027
27.7317 | -35.65 | -38.008
27.7378 | -35.6503 | -38.0131
27.7439 | -35.6506 | -38.0182
27.75 | -35.6509 | -38.0236
27.7561 | -35.6512 | -38.0286
27.7622 | -35.6514 | -38.0342
27.7683 | -35.6517 | -38.0391
27.7744 | -35.652 | -38.0442
27.7805 | -35.6523 | -38.05
27.7866 | -35.6526 | -38.055
27.7927 | -35.6528 | -38.0598
27.7988 | -35.6531 | -38.0654
27.8049 | -35.6534 | -38.0705
27.811 | -35.6535 | -38.0757
27.8171 | -35.6538 | -38.081
27.8232 | -35.6541 | -38.0858
27.8293 | -35.6544 | -38.0909
27.8354 | -35.6547 | -38.0962
27.8415 | -35.6549 | -38.1012
27.8476 | -35.6552 | -38.1064
27.8537 | -35.6555 | -38.1113
27.8598 | -35.6558 | -38.1162
27.866 | -35.656 | -38.1216
27.8721 | -35.6563 | -38.1263
27.8782 | -35.6566 | -38.1314
27.8843 | -35.6569 | -38.1363
27.8904 | -35.6572 | -38.1418
27.8965 | -35.6574 | -38.1467
27.9026 | -35.6577 | -38.1516
27.9087 | -35.658 | -38.1569
27.9148 | -35.6583 | -38.1623
27.9209 | -35.6586 | -38.1673
27.927 | -35.6588 | -38.1729
27.9331 | -35.6591 | -38.1777
27.9392 | -35.6594 | -38.1828
27.9453 | -35.6597 | -38.1883
27.9514 | -35.66 | -38.1942
27.9575 | -35.6602 | -38.1994
27.9636 | -35.6605 | -38.2045
27.9697 | -35.6608 | -38.2095
27.9758 | -35.661 | -38.215
27.9819 | -35.6612 | -38.2207
27.988 | -35.6615 | -38.2259
27.9941 | -35.6618 | -38.2314
28.0002 | -35.662 | -38.2365
28.0063 | -35.6622 | -38.2421
28.0125 | -35.6624 | -38.2471
28.0186 | -35.6628 | -38.2522
28.0247 | -35.663 | -38.2569
28.0308 | -35.6634 | -38.2621
28.0369 | -35.6636 | -38.2671
28.043 | -35.6639 | -38.2723
28.0491 | -35.6642 | -38.2773
28.0552 | -35.6644 | -38.2825
28.0613 | -35.6647 | -38.2878
28.0674 | -35.665 | -38.2928
28.0735 | -35.6652 | -38.2978

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.7112 | -43.1542
21.4908 | -35.4911 | -35.3018
16.4948 | -34.88 | -27.8606
12.6602 | -34.5228 | -21.5382
9.7171 | -34.3078 | -16.6058
7.4582 | -34.1644 | -12.601
5.7244 | -34.0657 | -9.3622
4.3936 | -34.02 | -6.743
3.3722 | -33.9911 | -4.6842
2.5883 | -33.9382 | -3.1201
1.9866 | -33.9381 | -1.8763
1.5248 | -33.9521 | -0.87492
1.1703 | -33.9669 | -0.22721
1 | -33.9556 | 0.09656
0.89822 | -33.9408 | 0.36572
0.68941 | -33.9267 | 0.94651
0.52919 | -33.922 | 1.3987
0.40612 | -33.9465 | 2.0069
0.31171 | -33.9482 | 2.4577
0.23928 | -33.9181 | 3.1871
0.18358 | -33.9277 | 3.6042
0.14091 | -33.9448 | 4.4832
0.10819 | -33.9816 | 5.6459
0.083017 | -33.9919 | 7.1311
0.063705 | -34.0338 | 9.0079
0.048923 | -34.1002 | 11.5134
0.037479 | -34.2064 | 14.7183
0.028801 | -34.3944 | 18.741
0.022078 | -34.6962 | 24.0631
0.017023 | -35.1644 | 30.4565
0.013018 | -35.9652 | 38.1437
0.010014 | -37.0752 | 46.065

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003689 | -8.365
0.007377 | -26.3493
0.011066 | -34.0698
0.014754 | -37.4571
0.018443 | -41.7092
0.022131 | -43.6647
0.02582 | -44.6172
0.029508 | -46.0352
0.033197 | -47.4126
0.036885 | -51.0957
0.040574 | -51.8595
0.044262 | -52.7495
0.047951 | -53.5921
0.051639 | -54.7206
0.055328 | -55.7197
0.059016 | -55.5071
0.062705 | -55.9373
0.066393 | -57.0427
0.070082 | -58.3784
0.07377 | -59.033
0.077459 | -58.7765
0.081147 | -59.4406
0.084836 | -60.3378
0.088524 | -60.8738
0.092213 | -61.5566
0.095901 | -61.656
0.09959 | -61.3856
0.10328 | -63.2627
0.10697 | -62.9338
0.11066 | -62.9954
0.11434 | -63.0053
0.11803 | -63.1444
0.12172 | -65.2189
0.12541 | -65.7329
0.1291 | -64.9913
0.13279 | -65.7287
0.13648 | -65.9027
0.14016 | -65.6202
0.14385 | -65.8212
0.14754 | -66.6104
0.15123 | -67.2617
0.15492 | -67.3806
0.15861 | -68.1389
0.16229 | -68.09
0.16598 | -68.1428
0.16967 | -68.367
0.17336 | -68.0713
0.17705 | -68.2779
0.18074 | -69.2953
0.18443 | -70.4276
0.18811 | -70.3046
0.1918 | -69.8349
0.19549 | -70.4908
0.19918 | -69.6423
0.20287 | -69.8857
0.20656 | -71.1622
0.21024 | -70.6364
0.21393 | -70.714
0.21762 | -72.0753
0.22131 | -71.6355
0.225 | -71.191
0.22869 | -70.8007
0.23238 | -71.8775
0.23606 | -72.1409
0.23975 | -72.1216
0.24344 | -72.5875
0.24713 | -72.1867
0.25082 | -72.626
0.25451 | -72.4135
0.2582 | -71.9823
0.26188 | -72.1829
0.26557 | -73.1915
0.26926 | -73.1136
0.27295 | -73.4545
0.27664 | -74.27
0.28033 | -74.6146
0.28402 | -73.9733
0.2877 | -74.5323
0.29139 | -73.5978
0.29508 | -73.9937
0.29877 | -74.3162
0.30246 | -74.7197
0.30615 | -74.3365
0.30983 | -74.9503
0.31352 | -75.177
0.31721 | -74.2579
0.3209 | -74.5358
0.32459 | -75.3784
0.32828 | -75.1381
0.33197 | -75.1085
0.33565 | -75.4853
0.33934 | -75.7982
0.34303 | -74.913
0.34672 | -75.6636
0.35041 | -75.9899
0.3541 | -75.5913
0.35779 | -76.1448
0.36147 | -76.0286
0.36516 | -75.7404
0.36885 | -75.5113
0.37254 | -75.7099
0.37623 | -76.169
0.37992 | -76.397
0.3836 | -76.8622
0.38729 | -76.3526
0.39098 | -76.3838
0.39467 | -77.1339
0.39836 | -77.2578
0.40205 | -76.7545
0.40574 | -77.9512
0.40942 | -77.9144
0.41311 | -76.9375
0.4168 | -77.6304
0.42049 | -78.124
0.42418 | -77.8056
0.42787 | -79.129
0.43156 | -77.7152
0.43524 | -77.1883
0.43893 | -77.4412
0.44262 | -78.1679
0.44631 | -77.3194
0.45 | -77.6912
0.45369 | -78.1685
0.45737 | -78.9759
0.46106 | -79.0092
0.46475 | -79.1256
0.46844 | -79.1167
0.47213 | -78.6815
0.47582 | -79.6144
0.47951 | -80.0704
0.48319 | -80.1474
0.48688 | -80.0372
0.49057 | -79.7742
0.49426 | -79.7362
0.49795 | -80.0436
0.50164 | -79.8379
0.50533 | -79.6137
0.50901 | -79.9994
0.5127 | -80.0664
0.51639 | -80.7902
0.52008 | -79.6833
0.52377 | -78.9911
0.52746 | -78.9669
0.53114 | -79.6016
0.53483 | -79.9555
0.53852 | -80.3392
0.54221 | -80.6793
0.5459 | -80.5448
0.54959 | -80.236
0.55328 | -80.2672
0.55696 | -80.7012
0.56065 | -81.1907
0.56434 | -81.5966
0.56803 | -81.7096
0.57172 | -81.3277
0.57541 | -80.4089
0.5791 | -80.0993
0.58278 | -80.2095
0.58647 | -80.1193
0.59016 | -80.5243
0.59385 | -80.9044
0.59754 | -80.9838
0.60123 | -81.2258
0.60491 | -81.2819
0.6086 | -81.1971
0.61229 | -81.3136
0.61598 | -81.3065
0.61967 | -81.3753
0.62336 | -81.7281
0.62705 | -81.7827
0.63073 | -81.4539
0.63442 | -81.4548
0.63811 | -81.3298
0.6418 | -80.9485
0.64549 | -81.1674
0.64918 | -81.9551
0.65287 | -82.1787
0.65655 | -81.7207
0.66024 | -81.5348
0.66393 | -81.7592
0.66762 | -82.0596
0.67131 | -82.1877
0.675 | -81.9956
0.67868 | -81.8828
0.68237 | -81.91
0.68606 | -82.2078
0.68975 | -82.6014
0.69344 | -83.2305
0.69713 | -83.7768
0.70082 | -83.5542
0.70451 | -83.3763
0.70819 | -83.5609
0.71188 | -83.5829
0.71557 | -83.082
0.71926 | -82.3533
0.72295 | -81.9172
0.72664 | -82.2748
0.73032 | -82.9318
0.73401 | -83.0247
0.7377 | -82.6761
0.74139 | -82.4606
0.74508 | -82.5914
0.74877 | -82.7487
0.75245 | -82.8967
0.75614 | -83.0025
0.75983 | -83.6038
0.76352 | -84.3039
0.76721 | -83.8058
0.7709 | -82.8738
0.77459 | -82.5752
0.77828 | -82.6601
0.78196 | -83.1405
0.78565 | -83.7464
0.78934 | -83.8157
0.79303 | -83.6301
0.79672 | -83.9623
0.80041 | -84.155
0.80409 | -83.4732
0.80778 | -83.2382
0.81147 | -83.7202
0.81516 | -84.5957
0.81885 | -85.475
0.82254 | -85.5474
0.82622 | -84.9864
0.82991 | -84.4714
0.8336 | -84.4801
0.83729 | -84.5661
0.84098 | -84.3925
0.84467 | -84.2113
0.84836 | -84.6488
0.85205 | -85.2897
0.85573 | -84.9816
0.85942 | -84.6757
0.86311 | -84.7952
0.8668 | -84.8939
0.87049 | -85.1585
0.87418 | -85.1946
0.87786 | -84.7353
0.88155 | -84.7343
0.88524 | -85.1651
0.88893 | -84.9389
0.89262 | -84.3094
0.89631 | -84.1892
0.89999 | -84.4375
0.90368 | -84.5529
0.90737 | -84.7004
0.91106 | -84.802
0.91475 | -84.7743
0.91844 | -85.3557
0.92213 | -86.0594
0.92582 | -85.849
0.9295 | -85.8276
0.93319 | -86.6008
0.93688 | -86.6739
0.94057 | -85.6869
0.94426 | -85.4102
0.94795 | -85.8744
0.95163 | -86.2139
0.95532 | -85.9684
0.95901 | -85.5563
0.9627 | -85.7756
0.96639 | -86.3128
0.97008 | -86.2664
0.97377 | -85.9307
0.97745 | -85.5888
0.98114 | -85.5159
0.98483 | -85.7329
0.98852 | -85.4987
0.99221 | -85.4725
0.9959 | -85.7568
0.99958 | -85.9618
1.0033 | -86.1707
1.007 | -86.065
1.0107 | -85.7978
1.0143 | -85.9817
1.018 | -86.6462
1.0217 | -86.8673
1.0254 | -86.2927
1.0291 | -85.7518
1.0328 | -85.5053
1.0365 | -85.5528
1.0402 | -85.7468
1.0438 | -86.0645
1.0475 | -86.5381
1.0512 | -87.055
1.0549 | -87.2787
1.0586 | -86.7225
1.0623 | -86.2113
1.066 | -86.1505
1.0697 | -86.3565
1.0734 | -86.5391
1.077 | -86.6545
1.0807 | -86.7405
1.0844 | -86.8387
1.0881 | -86.9709
1.0918 | -87.0299
1.0955 | -87.0249
1.0992 | -87.1179
1.1029 | -87.2553
1.1066 | -87.2388
1.1102 | -87.137
1.1139 | -87.1753
1.1176 | -87.3763
1.1213 | -87.5133
1.125 | -87.403
1.1287 | -87.2063
1.1324 | -87.1286
1.1361 | -87.1775
1.1397 | -87.2438
1.1434 | -87.2075
1.1471 | -87.0656
1.1508 | -86.913
1.1545 | -86.8372
1.1582 | -86.8603
1.1619 | -86.927
1.1656 | -86.9805
1.1693 | -87.0564
1.1729 | -87.1963
1.1766 | -87.3466
1.1803 | -87.3674
1.184 | -87.1621
1.1877 | -86.8766
1.1914 | -86.7934
1.1951 | -86.9763
1.1988 | -87.2541
1.2025 | -87.3876
1.2061 | -87.3248
1.2098 | -87.2806
1.2135 | -87.3895
1.2172 | -87.5054
1.2209 | -87.5256
1.2246 | -87.6277
1.2283 | -87.8632
1.232 | -87.9504
1.2356 | -87.76
1.2393 | -87.532
1.243 | -87.4627
1.2467 | -87.5521
1.2504 | -87.7966
1.2541 | -88.1314
1.2578 | -88.3598
1.2615 | -88.4244
1.2652 | -88.4252
1.2688 | -88.5005
1.2725 | -88.6254
1.2762 | -88.6037
1.2799 | -88.3483
1.2836 | -88.1037
1.2873 | -88.128
1.291 | -88.218
1.2947 | -88.0164
1.2984 | -87.7566
1.302 | -87.809
1.3057 | -88.0353
1.3094 | -88.1123
1.3131 | -88.09
1.3168 | -88.1666
1.3205 | -88.3513
1.3242 | -88.4099
1.3279 | -88.2543
1.3316 | -88.2049
1.3352 | -88.4614
1.3389 | -88.8686
1.3426 | -89.1597
1.3463 | -89.2307
1.35 | -89.1184
1.3537 | -88.9384
1.3574 | -88.9479
1.3611 | -89.1988
1.3647 | -89.3823
1.3684 | -89.2778
1.3721 | -88.9261
1.3758 | -88.5434
1.3795 | -88.3774
1.3832 | -88.41
1.3869 | -88.5327
1.3906 | -88.7304
1.3943 | -88.9106
1.3979 | -88.9854
1.4016 | -89.0228
1.4053 | -89.0556
1.409 | -89.0706
1.4127 | -89.0935
1.4164 | -89.0358
1.4201 | -88.884
1.4238 | -88.8759
1.4275 | -89.1315
1.4311 | -89.4254
1.4348 | -89.4546
1.4385 | -89.3275
1.4422 | -89.326
1.4459 | -89.4236
1.4496 | -89.4745
1.4533 | -89.4571
1.457 | -89.4157
1.4606 | -89.4174
1.4643 | -89.5266
1.468 | -89.7318
1.4717 | -89.9215
1.4754 | -90.0213
1.4791 | -90.0155
1.4828 | -89.9607
1.4865 | -89.9195
1.4902 | -89.8014
1.4938 | -89.6517
1.4975 | -89.6449
1.5012 | -89.7097
1.5049 | -89.6752
1.5086 | -89.6196
1.5123 | -89.7947
1.516 | -90.1476
1.5197 | -90.2944
1.5234 | -90.1831
1.527 | -90.14
1.5307 | -90.3038
1.5344 | -90.4642
1.5381 | -90.4319
1.5418 | -90.3852
1.5455 | -90.4342
1.5492 | -90.4704
1.5529 | -90.4454
1.5565 | -90.3011
1.5602 | -89.901
1.5639 | -89.4215
1.5676 | -89.2468
1.5713 | -89.4627
1.575 | -89.7803
1.5787 | -90.0262
1.5824 | -90.4048
1.5861 | -90.9344
1.5897 | -91.2323
1.5934 | -91.1684
1.5971 | -91.0931
1.6008 | -90.8846
1.6045 | -90.5975
1.6082 | -90.3576
1.6119 | -90.2492
1.6156 | -90.2846
1.6193 | -90.4213
1.6229 | -90.5843
1.6266 | -90.7209
1.6303 | -90.7806
1.634 | -90.7428
1.6377 | -90.6316
1.6414 | -90.5178
1.6451 | -90.4851
1.6488 | -90.6095
1.6525 | -90.9
1.6561 | -91.2424
1.6598 | -91.4664
1.6635 | -91.5249
1.6672 | -91.4273
1.6709 | -91.2152
1.6746 | -90.9278
1.6783 | -90.6717
1.682 | -90.5232
1.6856 | -90.5241
1.6893 | -90.6524
1.693 | -90.8477
1.6967 | -90.9993
1.7004 | -90.9976
1.7041 | -90.7962
1.7078 | -90.487
1.7115 | -90.1972
1.7152 | -90.0333
1.7188 | -90.0584
1.7225 | -90.2956
1.7262 | -90.6693
1.7299 | -91.0151
1.7336 | -91.1918
1.7373 | -91.2108
1.741 | -91.157
1.7447 | -91.0989
1.7484 | -91.1125
1.752 | -91.2365
1.7557 | -91.4049
1.7594 | -91.5044
1.7631 | -91.434
1.7668 | -91.1679
1.7705 | -90.8598
1.7742 | -90.7139
1.7779 | -90.7876
1.7815 | -90.9795
1.7852 | -91.164
1.7889 | -91.25
1.7926 | -91.2046
1.7963 | -91.0848
1.8 | -90.9532
1.8037 | -90.8265
1.8074 | -90.7473
1.8111 | -90.7821
1.8147 | -90.9412
1.8184 | -91.1098
1.8221 | -91.1453
1.8258 | -91.0293
1.8295 | -90.8882
1.8332 | -90.8683
1.8369 | -91.0001
1.8406 | -91.2397
1.8443 | -91.4978
1.8479 | -91.6629
1.8516 | -91.673
1.8553 | -91.5514
1.859 | -91.4233
1.8627 | -91.394
1.8664 | -91.45
1.8701 | -91.5327
1.8738 | -91.5712
1.8774 | -91.5158
1.8811 | -91.3856
1.8848 | -91.2859
1.8885 | -91.3214
1.8922 | -91.4842
1.8959 | -91.7228
1.8996 | -91.9636
1.9033 | -92.1537
1.907 | -92.2578
1.9106 | -92.266
1.9143 | -92.194
1.918 | -92.1359
1.9217 | -92.1761
1.9254 | -92.2971
1.9291 | -92.408
1.9328 | -92.4722
1.9365 | -92.4261
1.9402 | -92.2545
1.9438 | -92.0368
1.9475 | -91.9145
1.9512 | -91.9573
1.9549 | -92.1226
1.9586 | -92.3119
1.9623 | -92.4372
1.966 | -92.477
1.9697 | -92.448
1.9734 | -92.3551
1.977 | -92.206
1.9807 | -92.0531
1.9844 | -92.0025
1.9881 | -92.0914
1.9918 | -92.2057
1.9955 | -92.2258
1.9992 | -92.1666
2.0029 | -92.1192
2.0065 | -92.1506
2.0102 | -92.2671
2.0139 | -92.4192
2.0176 | -92.5396
2.0213 | -92.5928
2.025 | -92.5755
2.0287 | -92.5028
2.0324 | -92.4094
2.0361 | -92.3257
2.0397 | -92.1793
2.0434 | -91.9452
2.0471 | -91.7288
2.0508 | -91.672
2.0545 | -91.81
2.0582 | -92.0211
2.0619 | -92.1273
2.0656 | -92.0669
2.0693 | -91.9737
2.0729 | -92.0166
2.0766 | -92.2046
2.0803 | -92.4479
2.084 | -92.6535
2.0877 | -92.7956
2.0914 | -92.8921
2.0951 | -92.9404
2.0988 | -92.921
2.1024 | -92.8204
2.1061 | -92.6315
2.1098 | -92.3947
2.1135 | -92.225
2.1172 | -92.2222
2.1209 | -92.3534
2.1246 | -92.4984
2.1283 | -92.4725
2.132 | -92.4573
2.1356 | -92.4379
2.1393 | -92.4489
2.143 | -92.5017
2.1467 | -92.5596
2.1504 | -92.556
2.1541 | -92.4734
2.1578 | -92.3726
2.1615 | -92.297
2.1652 | -92.2857
2.1688 | -92.3524
2.1725 | -92.4819
2.1762 | -92.6732
2.1799 | -92.9284
2.1836 | -93.187
2.1873 | -93.3535
2.191 | -93.3761
2.1947 | -93.2641
2.1983 | -93.0906
2.202 | -92.9662
2.2057 | -92.9227
2.2094 | -92.9675
2.2131 | -93.0689
2.2168 | -93.1891
2.2205 | -93.3151
2.2242 | -93.4317
2.2279 | -93.5244
2.2315 | -93.5385
2.2352 | -93.4696
2.2389 | -93.3188
2.2426 | -93.115
2.2463 | -92.8924
2.25 | -92.706
2.2537 | -92.6214
2.2574 | -92.6686
2.2611 | -92.8242
2.2647 | -93.0366
2.2684 | -93.2379
2.2721 | -93.3987
2.2758 | -93.4838
2.2795 | -93.475
2.2832 | -93.3832
2.2869 | -93.2924
2.2906 | -93.2576
2.2943 | -93.2854
2.2979 | -93.3415
2.3016 | -93.3752
2.3053 | -93.3852
2.309 | -93.404
2.3127 | -93.4548
2.3164 | -93.4964
2.3201 | -93.5147
2.3238 | -93.5211
2.3274 | -93.5284
2.3311 | -93.5539
2.3348 | -93.5579
2.3385 | -93.4873
2.3422 | -93.3583
2.3459 | -93.2538
2.3496 | -93.2616
2.3533 | -93.4091
2.357 | -93.6429
2.3606 | -93.8484
2.3643 | -93.9441
2.368 | -93.9081
2.3717 | -93.7462
2.3754 | -93.5311
2.3791 | -93.3422
2.3828 | -93.2446
2.3865 | -93.2211
2.3902 | -93.2208
2.3938 | -93.2164
2.3975 | -93.245
2.4012 | -93.3648
2.4049 | -93.5908
2.4086 | -93.8545
2.4123 | -94.0542
2.416 | -94.1286
2.4197 | -94.0875
2.4233 | -93.9931
2.427 | -93.8882
2.4307 | -93.8119
2.4344 | -93.7898
2.4381 | -93.8143
2.4418 | -93.8558
2.4455 | -93.8956
2.4492 | -93.9565
2.4529 | -94.0483
2.4565 | -94.1765
2.4602 | -94.341
2.4639 | -94.5267
2.4676 | -94.7195
2.4713 | -94.8388
2.475 | -94.8423
2.4787 | -94.7481
2.4824 | -94.5928
2.4861 | -94.3842
2.4897 | -94.1381
2.4934 | -93.9126
2.4971 | -93.7628
2.5008 | -93.7476
2.5045 | -93.8556
2.5082 | -94.0282
2.5119 | -94.1993
2.5156 | -94.3162
2.5192 | -94.3691
2.5229 | -94.3919
2.5266 | -94.4266
2.5303 | -94.4829
2.534 | -94.5246
2.5377 | -94.5256
2.5414 | -94.4677
2.5451 | -94.4186
2.5488 | -94.4421
2.5524 | -94.522
2.5561 | -94.6212
2.5598 | -94.7034
2.5635 | -94.7494
2.5672 | -94.7652
2.5709 | -94.7464
2.5746 | -94.694
2.5783 | -94.6082
2.582 | -94.5009
2.5856 | -94.3564
2.5893 | -94.186
2.593 | -94.0566
2.5967 | -94.0014
2.6004 | -94.0135
2.6041 | -94.042
2.6078 | -94.0446
2.6115 | -94.0187
2.6152 | -93.9866
2.6188 | -93.968
2.6225 | -94.0157
2.6262 | -94.1614
2.6299 | -94.4014
2.6336 | -94.6357
2.6373 | -94.7962
2.641 | -94.836
2.6447 | -94.7671
2.6483 | -94.6097
2.652 | -94.3939
2.6557 | -94.1537
2.6594 | -94.0422
2.6631 | -93.981
2.6668 | -94.0143
2.6705 | -94.1331
2.6742 | -94.2916
2.6779 | -94.4628
2.6815 | -94.5767
2.6852 | -94.6249
2.6889 | -94.6131
2.6926 | -94.5678
2.6963 | -94.5348
2.7 | -94.5478
2.7037 | -94.6153
2.7074 | -94.7253
2.7111 | -94.8301
2.7147 | -94.897
2.7184 | -94.9242
2.7221 | -94.915
2.7258 | -94.8871
2.7295 | -94.8777
2.7332 | -94.8953
2.7369 | -94.9351
2.7406 | -94.9988
2.7442 | -95.061
2.7479 | -95.1063
2.7516 | -95.1011
2.7553 | -95.0578
2.759 | -94.9929
2.7627 | -94.9202
2.7664 | -94.8425
2.7701 | -94.7501
2.7738 | -94.6353
2.7774 | -94.4883
2.7811 | -94.3435
2.7848 | -94.2421
2.7885 | -94.1957
2.7922 | -94.2152
2.7959 | -94.2798
2.7996 | -94.3535
2.8033 | -94.4046
2.807 | -94.451
2.8106 | -94.4865
2.8143 | -94.5089
2.818 | -94.5266
2.8217 | -94.541
2.8254 | -94.5545
2.8291 | -94.5575
2.8328 | -94.5579
2.8365 | -94.5603
2.8401 | -94.5688
2.8438 | -94.5883
2.8475 | -94.6189
2.8512 | -94.6617
2.8549 | -94.7164
2.8586 | -94.797
2.8623 | -94.9046
2.866 | -95.0301
2.8697 | -95.1518
2.8733 | -95.2794
2.877 | -95.3954
2.8807 | -95.4836
2.8844 | -95.529
2.8881 | -95.497
2.8918 | -95.3855
2.8955 | -95.2143
2.8992 | -95.0256
2.9029 | -94.8603
2.9065 | -94.7485
2.9102 | -94.677
2.9139 | -94.6381
2.9176 | -94.6335
2.9213 | -94.6474
2.925 | -94.6898
2.9287 | -94.762
2.9324 | -94.8622
2.9361 | -94.9966
2.9397 | -95.151
2.9434 | -95.3147
2.9471 | -95.469
2.9508 | -95.5991
2.9545 | -95.6998
2.9582 | -95.793
2.9619 | -95.8693
2.9656 | -95.9334
2.9692 | -95.9855
2.9729 | -96.0078
2.9766 | -96.0039
2.9803 | -95.9727
2.984 | -95.9312
2.9877 | -95.8905
2.9914 | -95.841
2.9951 | -95.7759
2.9988 | -95.7134
3.0024 | -95.6447
3.0061 | -95.5612
3.0098 | -95.4513
3.0135 | -95.3205
3.0172 | -95.1868
3.0209 | -95.0807
3.0246 | -95.0273
3.0283 | -95.044
3.032 | -95.1103
3.0356 | -95.2215
3.0393 | -95.3635
3.043 | -95.5361
3.0467 | -95.7188
3.0504 | -95.8783
3.0541 | -95.9859
3.0578 | -96.03
3.0615 | -95.9873
3.0651 | -95.8916
3.0688 | -95.7741
3.0725 | -95.6707
3.0762 | -95.6083
3.0799 | -95.6108
3.0836 | -95.6774
3.0873 | -95.7773
3.091 | -95.8715
3.0947 | -95.9398
3.0983 | -95.9704
3.102 | -95.9682
3.1057 | -95.9452
3.1094 | -95.9168
3.1131 | -95.8918
3.1168 | -95.8853
3.1205 | -95.9067
3.1242 | -95.9434
3.1279 | -95.9654
3.1315 | -95.961
3.1352 | -95.9238
3.1389 | -95.8596
3.1426 | -95.7861
3.1463 | -95.7456
3.15 | -95.735
3.1537 | -95.7666
3.1574 | -95.8424
3.161 | -95.9408
3.1647 | -96.0489
3.1684 | -96.1607
3.1721 | -96.263
3.1758 | -96.3583
3.1795 | -96.4325
3.1832 | -96.4862
3.1869 | -96.4937
3.1906 | -96.4521
3.1942 | -96.4168
3.1979 | -96.3457
3.2016 | -96.2566
3.2053 | -96.139
3.209 | -96.0123
3.2127 | -95.8924
3.2164 | -95.8092
3.2201 | -95.7782
3.2238 | -95.8145
3.2274 | -95.895
3.2311 | -96.0182
3.2348 | -96.1625
3.2385 | -96.3074
3.2422 | -96.4137
3.2459 | -96.47
3.2496 | -96.4775
3.2533 | -96.4599
3.257 | -96.4277
3.2606 | -96.4104
3.2643 | -96.4149
3.268 | -96.4526
3.2717 | -96.5136
3.2754 | -96.579
3.2791 | -96.6365
3.2828 | -96.6686
3.2865 | -96.6463
3.2901 | -96.5928
3.2938 | -96.5118
3.2975 | -96.4127
3.3012 | -96.2975
3.3049 | -96.1777
3.3086 | -96.0625
3.3123 | -95.9674
3.316 | -95.9012
3.3197 | -95.8707
3.3233 | -95.8787
3.327 | -95.9303
3.3307 | -96.0276
3.3344 | -96.1519
3.3381 | -96.2826
3.3418 | -96.3938
3.3455 | -96.4465
3.3492 | -96.4562
3.3529 | -96.4281
3.3565 | -96.3851
3.3602 | -96.3383
3.3639 | -96.2968
3.3676 | -96.2694
3.3713 | -96.2663
3.375 | -96.2906
3.3787 | -96.356
3.3824 | -96.461
3.386 | -96.6076
3.3897 | -96.7806
3.3934 | -96.9497
3.3971 | -97.0877
3.4008 | -97.1471
3.4045 | -97.1039
3.4082 | -96.989
3.4119 | -96.8282
3.4156 | -96.6568
3.4192 | -96.5183
3.4229 | -96.4238
3.4266 | -96.3862
3.4303 | -96.392
3.434 | -96.4269
3.4377 | -96.4778
3.4414 | -96.5209
3.4451 | -96.5466
3.4488 | -96.5651
3.4524 | -96.5727
3.4561 | -96.5813
3.4598 | -96.5977
3.4635 | -96.5996
3.4672 | -96.5937
3.4709 | -96.5718
3.4746 | -96.5287
3.4783 | -96.4658
3.4819 | -96.3873
3.4856 | -96.3195
3.4893 | -96.2676
3.493 | -96.2483
3.4967 | -96.2817
3.5004 | -96.3641
3.5041 | -96.4819
3.5078 | -96.6278
3.5115 | -96.7849
3.5151 | -96.9252
3.5188 | -97.0439
3.5225 | -97.1301
3.5262 | -97.1995
3.5299 | -97.2595
3.5336 | -97.3199
3.5373 | -97.3847
3.541 | -97.4319
3.5447 | -97.4503
3.5483 | -97.4146
3.552 | -97.3128
3.5557 | -97.1696
3.5594 | -97.0039
3.5631 | -96.8481
3.5668 | -96.7258
3.5705 | -96.6468
3.5742 | -96.6291
3.5779 | -96.6626
3.5815 | -96.7364
3.5852 | -96.8344
3.5889 | -96.9456
3.5926 | -97.0593
3.5963 | -97.1721
3.6 | -97.2678
3.6037 | -97.3517
3.6074 | -97.4065
3.611 | -97.419
3.6147 | -97.4105
3.6184 | -97.3777
3.6221 | -97.3359
3.6258 | -97.2966
3.6295 | -97.2628
3.6332 | -97.249
3.6369 | -97.2716
3.6406 | -97.3395
3.6442 | -97.4461
3.6479 | -97.5633
3.6516 | -97.6696
3.6553 | -97.7324
3.659 | -97.7224
3.6627 | -97.6458
3.6664 | -97.5202
3.6701 | -97.3636
3.6738 | -97.2046
3.6774 | -97.0475
3.6811 | -96.9175
3.6848 | -96.8358
3.6885 | -96.8163
3.6922 | -96.8545
3.6959 | -96.9345
3.6996 | -97.0245
3.7033 | -97.1186
3.7069 | -97.1911
3.7106 | -97.2389
3.7143 | -97.2553
3.718 | -97.2476
3.7217 | -97.2403
3.7254 | -97.2364
3.7291 | -97.2349
3.7328 | -97.2505
3.7365 | -97.2796
3.7401 | -97.3072
3.7438 | -97.3295
3.7475 | -97.345
3.7512 | -97.3607
3.7549 | -97.3817
3.7586 | -97.3975
3.7623 | -97.39
3.766 | -97.369
3.7697 | -97.3269
3.7733 | -97.2648
3.777 | -97.1794
3.7807 | -97.0823
3.7844 | -96.9974
3.7881 | -96.938
3.7918 | -96.9079
3.7955 | -96.9154
3.7992 | -96.9597
3.8028 | -97.0415
3.8065 | -97.1411
3.8102 | -97.2366
3.8139 | -97.3134
3.8176 | -97.3615
3.8213 | -97.3889
3.825 | -97.3958
3.8287 | -97.3912
3.8324 | -97.3931
3.836 | -97.4176
3.8397 | -97.4597
3.8434 | -97.5038
3.8471 | -97.542
3.8508 | -97.5582
3.8545 | -97.5507
3.8582 | -97.5183
3.8619 | -97.4684
3.8656 | -97.406
3.8692 | -97.359
3.8729 | -97.3234
3.8766 | -97.2987
3.8803 | -97.2872
3.884 | -97.2853
3.8877 | -97.297
3.8914 | -97.3157
3.8951 | -97.3404
3.8988 | -97.3788
3.9024 | -97.433
3.9061 | -97.4955
3.9098 | -97.5695
3.9135 | -97.6536
3.9172 | -97.7306
3.9209 | -97.7988
3.9246 | -97.8511
3.9283 | -97.8707
3.9319 | -97.8579
3.9356 | -97.8485
3.9393 | -97.8367
3.943 | -97.8267
3.9467 | -97.827
3.9504 | -97.8353
3.9541 | -97.8434
3.9578 | -97.8411
3.9615 | -97.8238
3.9651 | -97.7979
3.9688 | -97.7535
3.9725 | -97.6881
3.9762 | -97.619
3.9799 | -97.5611
3.9836 | -97.5197
3.9873 | -97.514
3.991 | -97.5471
3.9947 | -97.6138
3.9983 | -97.7053
4.002 | -97.8216
4.0057 | -97.9514
4.0094 | -98.0781
4.0131 | -98.1876
4.0168 | -98.2779
4.0205 | -98.3467
4.0242 | -98.3849
4.0278 | -98.3973
4.0315 | -98.4109
4.0352 | -98.4189
4.0389 | -98.42
4.0426 | -98.4237
4.0463 | -98.4242
4.05 | -98.4236
4.0537 | -98.4332
4.0574 | -98.4436
4.061 | -98.4553
4.0647 | -98.4623
4.0684 | -98.4664
4.0721 | -98.4737
4.0758 | -98.4898
4.0795 | -98.5125
4.0832 | -98.5364
4.0869 | -98.5659
4.0906 | -98.5883
4.0942 | -98.5793
4.0979 | -98.5497
4.1016 | -98.5038
4.1053 | -98.4427
4.109 | -98.3874
4.1127 | -98.3497
4.1164 | -98.3166
4.1201 | -98.2979
4.1237 | -98.2767
4.1274 | -98.2549
4.1311 | -98.2256
4.1348 | -98.1959
4.1385 | -98.1783
4.1422 | -98.1804
4.1459 | -98.207
4.1496 | -98.2477
4.1533 | -98.2938
4.1569 | -98.3324
4.1606 | -98.3516
4.1643 | -98.352
4.168 | -98.342
4.1717 | -98.3231
4.1754 | -98.3049
4.1791 | -98.3019
4.1828 | -98.3082
4.1865 | -98.3444
4.1901 | -98.4004
4.1938 | -98.468
4.1975 | -98.5362
4.2012 | -98.5932
4.2049 | -98.6352
4.2086 | -98.6483
4.2123 | -98.6422
4.216 | -98.6249
4.2197 | -98.5949
4.2233 | -98.5439
4.227 | -98.4829
4.2307 | -98.423
4.2344 | -98.3648
4.2381 | -98.3279
4.2418 | -98.3192
4.2455 | -98.3528
4.2492 | -98.4077
4.2528 | -98.4697
4.2565 | -98.4917
4.2602 | -98.5136
4.2639 | -98.5058
4.2676 | -98.4836
4.2713 | -98.4459
4.275 | -98.4092
4.2787 | -98.3718
4.2824 | -98.345
4.286 | -98.3312
4.2897 | -98.3284
4.2934 | -98.3385
4.2971 | -98.3695
4.3008 | -98.419
4.3045 | -98.4747
4.3082 | -98.532
4.3119 | -98.5749
4.3156 | -98.6008
4.3192 | -98.6031
4.3229 | -98.5717
4.3266 | -98.5068
4.3303 | -98.4354
4.334 | -98.3736
4.3377 | -98.3235
4.3414 | -98.304
4.3451 | -98.3126
4.3487 | -98.3593
4.3524 | -98.4351
4.3561 | -98.5383
4.3598 | -98.6545
4.3635 | -98.7671
4.3672 | -98.8619
4.3709 | -98.9311
4.3746 | -98.9643
4.3783 | -98.9621
4.3819 | -98.9264
4.3856 | -98.8739
4.3893 | -98.8211
4.393 | -98.7762
4.3967 | -98.7463
4.4004 | -98.7256
4.4041 | -98.7161
4.4078 | -98.7173
4.4115 | -98.7224
4.4151 | -98.7227
4.4188 | -98.7167
4.4225 | -98.7079
4.4262 | -98.7077
4.4299 | -98.707
4.4336 | -98.7154
4.4373 | -98.7267
4.441 | -98.7393
4.4446 | -98.7462
4.4483 | -98.748
4.452 | -98.7358
4.4557 | -98.7054
4.4594 | -98.6772
4.4631 | -98.6478
4.4668 | -98.6222
4.4705 | -98.6054
4.4742 | -98.6002
4.4778 | -98.6106
4.4815 | -98.6329
4.4852 | -98.6689
4.4889 | -98.7039
4.4926 | -98.7386
4.4963 | -98.772
4.5 | -98.821
4.5037 | -98.8626
4.5074 | -98.9107
4.511 | -98.9576
4.5147 | -98.996
4.5184 | -99.0298
4.5221 | -99.0586
4.5258 | -99.088
4.5295 | -99.0992
4.5332 | -99.1107
4.5369 | -99.1225
4.5406 | -99.1323
4.5442 | -99.1446
4.5479 | -99.1523
4.5516 | -99.1579
4.5553 | -99.1577
4.559 | -99.1541
4.5627 | -99.1478
4.5664 | -99.1403
4.5701 | -99.1269
4.5737 | -99.1246
4.5774 | -99.1193
4.5811 | -99.1131
4.5848 | -99.1017
4.5885 | -99.0775
4.5922 | -99.0386
4.5959 | -98.9834
4.5996 | -98.9252
4.6033 | -98.8602
4.6069 | -98.795
4.6106 | -98.7461
4.6143 | -98.7092
4.618 | -98.6871
4.6217 | -98.6815
4.6254 | -98.6927
4.6291 | -98.7158
4.6328 | -98.7467
4.6365 | -98.7851
4.6401 | -98.8327
4.6438 | -98.8793
4.6475 | -98.9293
4.6512 | -98.9736
4.6549 | -99.0037
4.6586 | -99.0144
4.6623 | -99.0217
4.666 | -99.0115
4.6696 | -98.9827
4.6733 | -98.9426
4.677 | -98.9035
4.6807 | -98.8663
4.6844 | -98.8505
4.6881 | -98.8378
4.6918 | -98.8377
4.6955 | -98.8505
4.6992 | -98.8701
4.7028 | -98.8959
4.7065 | -98.926
4.7102 | -98.9651
4.7139 | -99.0163
4.7176 | -99.0757
4.7213 | -99.1303
4.725 | -99.175
4.7287 | -99.2009
4.7324 | -99.1983
4.736 | -99.1736
4.7397 | -99.1323
4.7434 | -99.0842
4.7471 | -99.0354
4.7508 | -99.0058
4.7545 | -99.0016
4.7582 | -99.0363
4.7619 | -99.0845
4.7655 | -99.1486
4.7692 | -99.2233
4.7729 | -99.2943
4.7766 | -99.3575
4.7803 | -99.4013
4.784 | -99.4289
4.7877 | -99.4278
4.7914 | -99.4449
4.7951 | -99.4585
4.7987 | -99.4608
4.8024 | -99.4593
4.8061 | -99.456
4.8098 | -99.4604
4.8135 | -99.4746
4.8172 | -99.4945
4.8209 | -99.5207
4.8246 | -99.5467
4.8283 | -99.5657
4.8319 | -99.5798
4.8356 | -99.5908
4.8393 | -99.5852
4.843 | -99.5706
4.8467 | -99.5525
4.8504 | -99.5268
4.8541 | -99.5003
4.8578 | -99.475
4.8615 | -99.4625
4.8651 | -99.4546
4.8688 | -99.4633
4.8725 | -99.4881
4.8762 | -99.5237
4.8799 | -99.5601
4.8836 | -99.5901
4.8873 | -99.615
4.891 | -99.632
4.8946 | -99.6395
4.8983 | -99.6336
4.902 | -99.6096
4.9057 | -99.5787
4.9094 | -99.5476
4.9131 | -99.515
4.9168 | -99.4872
4.9205 | -99.4692
4.9242 | -99.4562
4.9278 | -99.4509
4.9315 | -99.4509
4.9352 | -99.4428
4.9389 | -99.4207
4.9426 | -99.3923
4.9463 | -99.35
4.95 | -99.3055
4.9537 | -99.2634
4.9574 | -99.2348
4.961 | -99.2308
4.9647 | -99.2459
4.9684 | -99.2848
4.9721 | -99.3465
4.9758 | -99.4258
4.9795 | -99.5165
4.9832 | -99.6081
4.9869 | -99.6946
4.9905 | -99.7671
4.9942 | -99.8229
4.9979 | -99.8516
5.0016 | -99.8487
5.0053 | -99.8198
5.009 | -99.7701
5.0127 | -99.7153
5.0164 | -99.6609
5.0201 | -99.6138
5.0237 | -99.5845
5.0274 | -99.5787
5.0311 | -99.5958
5.0348 | -99.6231
5.0385 | -99.6576
5.0422 | -99.7046
5.0459 | -99.7581
5.0496 | -99.8114
5.0533 | -99.8628
5.0569 | -99.906
5.0606 | -99.93
5.0643 | -99.9362
5.068 | -99.9255
5.0717 | -99.8912
5.0754 | -99.8455
5.0791 | -99.7949
5.0828 | -99.7321
5.0864 | -99.6625
5.0901 | -99.5952
5.0938 | -99.5342
5.0975 | -99.4741
5.1012 | -99.4218
5.1049 | -99.3789
5.1086 | -99.3505
5.1123 | -99.3443
5.116 | -99.3623
5.1196 | -99.3981
5.1233 | -99.4467
5.127 | -99.5142
5.1307 | -99.5996
5.1344 | -99.6992
5.1381 | -99.8054
5.1418 | -99.9141
5.1455 | -100.034
5.1492 | -100.1587
5.1528 | -100.273
5.1565 | -100.3833
5.1602 | -100.4995
5.1639 | -100.6147
5.1676 | -100.7137
5.1713 | -100.7954
5.175 | -100.8619
5.1787 | -100.8895
5.1824 | -100.8728
5.186 | -100.8237
5.1897 | -100.7324
5.1934 | -100.6048
5.1971 | -100.4594
5.2008 | -100.3042
5.2045 | -100.1457
5.2082 | -99.9959
5.2119 | -99.874
5.2155 | -99.7841
5.2192 | -99.729
5.2229 | -99.7094
5.2266 | -99.7219
5.2303 | -99.7638
5.234 | -99.8169
5.2377 | -99.8678
5.2414 | -99.9222
5.2451 | -99.9711
5.2487 | -100.0046
5.2524 | -100.0255
5.2561 | -100.0314
5.2598 | -100.0239
5.2635 | -100.0084
5.2672 | -99.9898
5.2709 | -99.9664
5.2746 | -99.9446
5.2783 | -99.9268
5.2819 | -99.9116
5.2856 | -99.897
5.2893 | -99.8891
5.293 | -99.8867
5.2967 | -99.8821
5.3004 | -99.8841
5.3041 | -99.8938
5.3078 | -99.9047
5.3114 | -99.9183
5.3151 | -99.9317
5.3188 | -99.9519
5.3225 | -99.9733
5.3262 | -99.9945
5.3299 | -100.0186
5.3336 | -100.0465
5.3373 | -100.0804
5.341 | -100.1118
5.3446 | -100.1435
5.3483 | -100.1781
5.352 | -100.2094
5.3557 | -100.2322
5.3594 | -100.2536
5.3631 | -100.2677
5.3668 | -100.2754
5.3705 | -100.2702
5.3742 | -100.2597
5.3778 | -100.2499
5.3815 | -100.2293
5.3852 | -100.2007
5.3889 | -100.1717
5.3926 | -100.1495
5.3963 | -100.1308
5.4 | -100.1218
5.4037 | -100.1192
5.4073 | -100.1192
5.411 | -100.1385
5.4147 | -100.1684
5.4184 | -100.2018
5.4221 | -100.2349
5.4258 | -100.2742
5.4295 | -100.3059
5.4332 | -100.3355
5.4369 | -100.3592
5.4405 | -100.3785
5.4442 | -100.3961
5.4479 | -100.4108
5.4516 | -100.4291
5.4553 | -100.4439
5.459 | -100.442
5.4627 | -100.4342
5.4664 | -100.4289
5.4701 | -100.4199
5.4737 | -100.4056
5.4774 | -100.3928
5.4811 | -100.3839
5.4848 | -100.3748
5.4885 | -100.3748
5.4922 | -100.3874
5.4959 | -100.3964
5.4996 | -100.404
5.5033 | -100.413
5.5069 | -100.4256
5.5106 | -100.4274
5.5143 | -100.4303
5.518 | -100.4251
5.5217 | -100.4233
5.5254 | -100.4207
5.5291 | -100.4178
5.5328 | -100.4258
5.5364 | -100.4338
5.5401 | -100.4525
5.5438 | -100.4771
5.5475 | -100.4913
5.5512 | -100.5
5.5549 | -100.5039
5.5586 | -100.5046
5.5623 | -100.4966
5.566 | -100.4879
5.5696 | -100.473
5.5733 | -100.453
5.577 | -100.4409
5.5807 | -100.4361
5.5844 | -100.4323
5.5881 | -100.4287
5.5918 | -100.4312
5.5955 | -100.4396
5.5992 | -100.4486
5.6028 | -100.4631
5.6065 | -100.4716
5.6102 | -100.4804
5.6139 | -100.4879
5.6176 | -100.4912
5.6213 | -100.4906
5.625 | -100.4851
5.6287 | -100.4758
5.6323 | -100.465
5.636 | -100.4525
5.6397 | -100.4372
5.6434 | -100.4132
5.6471 | -100.3803
5.6508 | -100.3463
5.6545 | -100.3167
5.6582 | -100.2866
5.6619 | -100.2519
5.6655 | -100.2204
5.6692 | -100.1951
5.6729 | -100.1799
5.6766 | -100.1661
5.6803 | -100.1533
5.684 | -100.1351
5.6877 | -100.1216
5.6914 | -100.1138
5.6951 | -100.1091
5.6987 | -100.1162
5.7024 | -100.1341
5.7061 | -100.1683
5.7098 | -100.211
5.7135 | -100.2601
5.7172 | -100.307
5.7209 | -100.3474
5.7246 | -100.3834
5.7282 | -100.4092
5.7319 | -100.4322
5.7356 | -100.4418
5.7393 | -100.4362
5.743 | -100.4209
5.7467 | -100.397
5.7504 | -100.373
5.7541 | -100.3513
5.7578 | -100.3316
5.7614 | -100.328
5.7651 | -100.3404
5.7688 | -100.3715
5.7725 | -100.411
5.7762 | -100.4585
5.7799 | -100.5112
5.7836 | -100.5591
5.7873 | -100.6066
5.791 | -100.6445
5.7946 | -100.6764
5.7983 | -100.6995
5.802 | -100.7218
5.8057 | -100.7382
5.8094 | -100.7452
5.8131 | -100.7402
5.8168 | -100.7146
5.8205 | -100.6824
5.8242 | -100.6502
5.8278 | -100.6219
5.8315 | -100.5933
5.8352 | -100.5673
5.8389 | -100.5507
5.8426 | -100.5366
5.8463 | -100.5253
5.85 | -100.5317
5.8537 | -100.5305
5.8573 | -100.5293
5.861 | -100.5249
5.8647 | -100.5235
5.8684 | -100.5271
5.8721 | -100.5325
5.8758 | -100.5328
5.8795 | -100.5298
5.8832 | -100.5215
5.8869 | -100.5129
5.8905 | -100.4963
5.8942 | -100.4785
5.8979 | -100.4657
5.9016 | -100.4576
5.9053 | -100.4506
5.909 | -100.4478
5.9127 | -100.4431
5.9164 | -100.4398
5.9201 | -100.4386
5.9237 | -100.4331
5.9274 | -100.4235
5.9311 | -100.4189
5.9348 | -100.404
5.9385 | -100.3848
5.9422 | -100.3605
5.9459 | -100.3389
5.9496 | -100.321
5.9532 | -100.3027
5.9569 | -100.2728
5.9606 | -100.2336
5.9643 | -100.2056
5.968 | -100.186
5.9717 | -100.1696
5.9754 | -100.162
5.9791 | -100.1633
5.9828 | -100.182
5.9864 | -100.2177
5.9901 | -100.2647
5.9938 | -100.3253
5.9975 | -100.3987
6.0012 | -100.4708
6.0049 | -100.5473
6.0086 | -100.6209
6.0123 | -100.6904
6.016 | -100.7565
6.0196 | -100.8093
6.0233 | -100.841
6.027 | -100.8673
6.0307 | -100.8795
6.0344 | -100.8789
6.0381 | -100.8671
6.0418 | -100.8508
6.0455 | -100.8423
6.0491 | -100.843
6.0528 | -100.8532
6.0565 | -100.8576
6.0602 | -100.8668
6.0639 | -100.8865
6.0676 | -100.9059
6.0713 | -100.9293
6.075 | -100.9529
6.0787 | -100.9733
6.0823 | -100.9935
6.086 | -101.0133
6.0897 | -101.0339
6.0934 | -101.0576
6.0971 | -101.0821
6.1008 | -101.1024
6.1045 | -101.1176
6.1082 | -101.1319
6.1119 | -101.1461
6.1155 | -101.1602
6.1192 | -101.175
6.1229 | -101.1842
6.1266 | -101.2001
6.1303 | -101.218
6.134 | -101.2371
6.1377 | -101.2535
6.1414 | -101.2744
6.1451 | -101.2946
6.1487 | -101.3162
6.1524 | -101.3398
6.1561 | -101.361
6.1598 | -101.383
6.1635 | -101.402
6.1672 | -101.4206
6.1709 | -101.4337
6.1746 | -101.4398
6.1782 | -101.4383
6.1819 | -101.4238
6.1856 | -101.3984
6.1893 | -101.3624
6.193 | -101.3228
6.1967 | -101.2847
6.2004 | -101.2448
6.2041 | -101.2096
6.2078 | -101.1908
6.2114 | -101.1865
6.2151 | -101.1989
6.2188 | -101.2223
6.2225 | -101.252
6.2262 | -101.2877
6.2299 | -101.3304
6.2336 | -101.3713
6.2373 | -101.4052
6.241 | -101.4275
6.2446 | -101.4363
6.2483 | -101.4367
6.252 | -101.4189
6.2557 | -101.3868
6.2594 | -101.346
6.2631 | -101.306
6.2668 | -101.2548
6.2705 | -101.1912
6.2741 | -101.1301
6.2778 | -101.0732
6.2815 | -101.02
6.2852 | -100.9689
6.2889 | -100.9248
6.2926 | -100.8947
6.2963 | -100.8818
6.3 | -100.8812
6.3037 | -100.894
6.3073 | -100.9215
6.311 | -100.9643
6.3147 | -101.0282
6.3184 | -101.1002
6.3221 | -101.1763
6.3258 | -101.2556
6.3295 | -101.3377
6.3332 | -101.4183
6.3369 | -101.4854
6.3405 | -101.5412
6.3442 | -101.591
6.3479 | -101.6274
6.3516 | -101.6483
6.3553 | -101.6565
6.359 | -101.6607
6.3627 | -101.6569
6.3664 | -101.6354
6.37 | -101.6033
6.3737 | -101.5682
6.3774 | -101.531
6.3811 | -101.4894
6.3848 | -101.4529
6.3885 | -101.4211
6.3922 | -101.3894
6.3959 | -101.3631
6.3996 | -101.3448
6.4032 | -101.3295
6.4069 | -101.3174
6.4106 | -101.3025
6.4143 | -101.2911
6.418 | -101.2804
6.4217 | -101.2719
6.4254 | -101.2704
6.4291 | -101.2731
6.4328 | -101.2795
6.4364 | -101.2879
6.4401 | -101.3044
6.4438 | -101.3299
6.4475 | -101.3594
6.4512 | -101.3901
6.4549 | -101.4218
6.4586 | -101.4512
6.4623 | -101.4809
6.466 | -101.5095
6.4696 | -101.5357
6.4733 | -101.5605
6.477 | -101.5831
6.4807 | -101.5953
6.4844 | -101.5917
6.4881 | -101.5778
6.4918 | -101.564
6.4955 | -101.5381
6.4991 | -101.5087
6.5028 | -101.4792
6.5065 | -101.4427
6.5102 | -101.4035
6.5139 | -101.3631
6.5176 | -101.3236
6.5213 | -101.2844
6.525 | -101.247
6.5287 | -101.2217
6.5323 | -101.2109
6.536 | -101.2106
6.5397 | -101.2207
6.5434 | -101.2411
6.5471 | -101.2713
6.5508 | -101.31
6.5545 | -101.3543
6.5582 | -101.3985
6.5619 | -101.4354
6.5655 | -101.4679
6.5692 | -101.4782
6.5729 | -101.4817
6.5766 | -101.4738
6.5803 | -101.4558
6.584 | -101.4272
6.5877 | -101.3936
6.5914 | -101.354
6.595 | -101.3197
6.5987 | -101.2893
6.6024 | -101.2706
6.6061 | -101.2658
6.6098 | -101.2763
6.6135 | -101.3012
6.6172 | -101.3376
6.6209 | -101.3818
6.6246 | -101.4298
6.6282 | -101.4768
6.6319 | -101.5194
6.6356 | -101.5526
6.6393 | -101.5784
6.643 | -101.5993
6.6467 | -101.6175
6.6504 | -101.6196
6.6541 | -101.6195
6.6578 | -101.624
6.6614 | -101.6258
6.6651 | -101.6228
6.6688 | -101.6341
6.6725 | -101.6385
6.6762 | -101.6457
6.6799 | -101.6569
6.6836 | -101.6703
6.6873 | -101.676
6.6909 | -101.6712
6.6946 | -101.6581
6.6983 | -101.6343
6.702 | -101.6062
6.7057 | -101.5733
6.7094 | -101.5403
6.7131 | -101.506
6.7168 | -101.4713
6.7205 | -101.4419
6.7241 | -101.4207
6.7278 | -101.4133
6.7315 | -101.4163
6.7352 | -101.4321
6.7389 | -101.4651
6.7426 | -101.5151
6.7463 | -101.5753
6.75 | -101.6481
6.7537 | -101.7205
6.7573 | -101.7851
6.761 | -101.8457
6.7647 | -101.8994
6.7684 | -101.9403
6.7721 | -101.9705
6.7758 | -101.9872
6.7795 | -101.9861
6.7832 | -101.9729
6.7869 | -101.9554
6.7905 | -101.9398
6.7942 | -101.9177
6.7979 | -101.8957
6.8016 | -101.8741
6.8053 | -101.8594
6.809 | -101.8463
6.8127 | -101.838
6.8164 | -101.838
6.82 | -101.8437
6.8237 | -101.8545
6.8274 | -101.8665
6.8311 | -101.8913
6.8348 | -101.9252
6.8385 | -101.9554
6.8422 | -101.9868
6.8459 | -102.0142
6.8496 | -102.0391
6.8532 | -102.0504
6.8569 | -102.053
6.8606 | -102.0406
6.8643 | -102.0166
6.868 | -101.9846
6.8717 | -101.943
6.8754 | -101.894
6.8791 | -101.8541
6.8828 | -101.8172
6.8864 | -101.7861
6.8901 | -101.7656
6.8938 | -101.7607
6.8975 | -101.76
6.9012 | -101.7715
6.9049 | -101.798
6.9086 | -101.8397
6.9123 | -101.8875
6.9159 | -101.9481
6.9196 | -102.0025
6.9233 | -102.0587
6.927 | -102.1124
6.9307 | -102.1617
6.9344 | -102.2003
6.9381 | -102.2279
6.9418 | -102.2504
6.9455 | -102.2637
6.9491 | -102.2703
6.9528 | -102.2624
6.9565 | -102.2477
6.9602 | -102.2241
6.9639 | -102.1954
6.9676 | -102.1595
6.9713 | -102.1211
6.975 | -102.084
6.9787 | -102.0498
6.9823 | -102.0099
6.986 | -101.9698
6.9897 | -101.9297
6.9934 | -101.8934
6.9971 | -101.8605
7.0008 | -101.8308
7.0045 | -101.8038
7.0082 | -101.7831
7.0118 | -101.7679
7.0155 | -101.7616
7.0192 | -101.7574
7.0229 | -101.7568
7.0266 | -101.7635
7.0303 | -101.7776
7.034 | -101.8002
7.0377 | -101.8316
7.0414 | -101.8675
7.045 | -101.9098
7.0487 | -101.9541
7.0524 | -101.9956
7.0561 | -102.0449
7.0598 | -102.098
7.0635 | -102.1464
7.0672 | -102.1912
7.0709 | -102.2289
7.0746 | -102.2633
7.0782 | -102.2909
7.0819 | -102.3074
7.0856 | -102.3099
7.0893 | -102.299
7.093 | -102.2806
7.0967 | -102.2552
7.1004 | -102.22
7.1041 | -102.1804
7.1078 | -102.1364
7.1114 | -102.0879
7.1151 | -102.0388
7.1188 | -101.9976
7.1225 | -101.9625
7.1262 | -101.9352
7.1299 | -101.9213
7.1336 | -101.9185
7.1373 | -101.9215
7.1409 | -101.9344
7.1446 | -101.9579
7.1483 | -101.9847
7.152 | -102.0068
7.1557 | -102.0275
7.1594 | -102.0477
7.1631 | -102.0568
7.1668 | -102.0557
7.1705 | -102.0495
7.1741 | -102.038
7.1778 | -102.0276
7.1815 | -102.0121
7.1852 | -102
7.1889 | -101.9932
7.1926 | -101.9938
7.1963 | -102.0048
7.2 | -102.019
7.2037 | -102.0428
7.2073 | -102.0728
7.211 | -102.1116
7.2147 | -102.159
7.2184 | -102.2084
7.2221 | -102.2584
7.2258 | -102.3112
7.2295 | -102.3626
7.2332 | -102.4096
7.2368 | -102.4455
7.2405 | -102.4743
7.2442 | -102.4959
7.2479 | -102.5041
7.2516 | -102.5014
7.2553 | -102.5009
7.259 | -102.4987
7.2627 | -102.4892
7.2664 | -102.4826
7.27 | -102.4805
7.2737 | -102.4806
7.2774 | -102.4832
7.2811 | -102.4916
7.2848 | -102.5058
7.2885 | -102.5269
7.2922 | -102.5513
7.2959 | -102.5795
7.2996 | -102.6
7.3032 | -102.6158
7.3069 | -102.6304
7.3106 | -102.641
7.3143 | -102.65
7.318 | -102.6538
7.3217 | -102.6525
7.3254 | -102.6419
7.3291 | -102.6278
7.3327 | -102.6093
7.3364 | -102.5929
7.3401 | -102.5787
7.3438 | -102.5623
7.3475 | -102.5515
7.3512 | -102.5469
7.3549 | -102.5439
7.3586 | -102.5388
7.3623 | -102.5378
7.3659 | -102.5454
7.3696 | -102.5564
7.3733 | -102.5648
7.377 | -102.5723
7.3807 | -102.5811
7.3844 | -102.5885
7.3881 | -102.5879
7.3918 | -102.58
7.3955 | -102.5694
7.3991 | -102.5542
7.4028 | -102.5377
7.4065 | -102.525
7.4102 | -102.5072
7.4139 | -102.4833
7.4176 | -102.4666
7.4213 | -102.4611
7.425 | -102.4614
7.4287 | -102.4664
7.4323 | -102.4869
7.436 | -102.5097
7.4397 | -102.5349
7.4434 | -102.566
7.4471 | -102.6059
7.4508 | -102.639
7.4545 | -102.6722
7.4582 | -102.7047
7.4618 | -102.7346
7.4655 | -102.7602
7.4692 | -102.7826
7.4729 | -102.7935
7.4766 | -102.7928
7.4803 | -102.7939
7.484 | -102.7886
7.4877 | -102.7804
7.4914 | -102.766
7.495 | -102.7535
7.4987 | -102.737
7.5024 | -102.7176
7.5061 | -102.6998
7.5098 | -102.6851
7.5135 | -102.663
7.5172 | -102.6432
7.5209 | -102.6278
7.5246 | -102.6205
7.5282 | -102.6112
7.5319 | -102.6048
7.5356 | -102.6072
7.5393 | -102.6132
7.543 | -102.615
7.5467 | -102.6235
7.5504 | -102.6346
7.5541 | -102.6448
7.5577 | -102.6515
7.5614 | -102.6544
7.5651 | -102.649
7.5688 | -102.6417
7.5725 | -102.6257
7.5762 | -102.6037
7.5799 | -102.5812
7.5836 | -102.5532
7.5873 | -102.522
7.5909 | -102.4877
7.5946 | -102.4452
7.5983 | -102.4055
7.602 | -102.3649
7.6057 | -102.3289
7.6094 | -102.2896
7.6131 | -102.2565
7.6168 | -102.2295
7.6205 | -102.2082
7.6241 | -102.1943
7.6278 | -102.1923
7.6315 | -102.1963
7.6352 | -102.2028
7.6389 | -102.2074
7.6426 | -102.2149
7.6463 | -102.2279
7.65 | -102.2424
7.6536 | -102.2573
7.6573 | -102.2774
7.661 | -102.2986
7.6647 | -102.3215
7.6684 | -102.3473
7.6721 | -102.3724
7.6758 | -102.3966
7.6795 | -102.4292
7.6832 | -102.4602
7.6868 | -102.4915
7.6905 | -102.5254
7.6942 | -102.5538
7.6979 | -102.5842
7.7016 | -102.6088
7.7053 | -102.6314
7.709 | -102.6518
7.7127 | -102.6646
7.7164 | -102.6761
7.72 | -102.6831
7.7237 | -102.6891
7.7274 | -102.6956
7.7311 | -102.6985
7.7348 | -102.6998
7.7385 | -102.6994
7.7422 | -102.7047
7.7459 | -102.7128
7.7496 | -102.7253
7.7532 | -102.7462
7.7569 | -102.772
7.7606 | -102.7975
7.7643 | -102.8291
7.768 | -102.8634
7.7717 | -102.8942
7.7754 | -102.9257
7.7791 | -102.9569
7.7827 | -102.9828
7.7864 | -103.0089
7.7901 | -103.0324
7.7938 | -103.0454
7.7975 | -103.0522
7.8012 | -103.0535
7.8049 | -103.0475
7.8086 | -103.0337
7.8123 | -103.013
7.8159 | -102.9895
7.8196 | -102.9616
7.8233 | -102.9278
7.827 | -102.8904
7.8307 | -102.8458
7.8344 | -102.7985
7.8381 | -102.7536
7.8418 | -102.7133
7.8455 | -102.678
7.8491 | -102.6513
7.8528 | -102.6346
7.8565 | -102.6207
7.8602 | -102.6157
7.8639 | -102.625
7.8676 | -102.6437
7.8713 | -102.6662
7.875 | -102.6975
7.8786 | -102.7351
7.8823 | -102.7771
7.886 | -102.818
7.8897 | -102.8598
7.8934 | -102.9036
7.8971 | -102.9498
7.9008 | -102.9926
7.9045 | -103.0315
7.9082 | -103.0647
7.9118 | -103.0974
7.9155 | -103.121
7.9192 | -103.1358
7.9229 | -103.1451
7.9266 | -103.1481
7.9303 | -103.1471
7.934 | -103.1337
7.9377 | -103.1116
7.9414 | -103.0875
7.945 | -103.0582
7.9487 | -103.0253
7.9524 | -102.9846
7.9561 | -102.9356
7.9598 | -102.8906
7.9635 | -102.8542
7.9672 | -102.8226
7.9709 | -102.7985
7.9745 | -102.7801
7.9782 | -102.7649
7.9819 | -102.7712
7.9856 | -102.7607
7.9893 | -102.7563
7.993 | -102.7523
7.9967 | -102.7492
8.0004 | -102.7573
8.0041 | -102.7674
8.0077 | -102.7745
8.0114 | -102.7844
8.0151 | -102.7916
8.0188 | -102.7989
8.0225 | -102.8066
8.0262 | -102.8175
8.0299 | -102.8224
8.0336 | -102.8223
8.0373 | -102.8232
8.0409 | -102.8184
8.0446 | -102.8068
8.0483 | -102.7931
8.052 | -102.7823
8.0557 | -102.771
8.0594 | -102.7607
8.0631 | -102.7525
8.0668 | -102.7441
8.0704 | -102.7382
8.0741 | -102.735
8.0778 | -102.7306
8.0815 | -102.723
8.0852 | -102.7204
8.0889 | -102.7188
8.0926 | -102.7192
8.0963 | -102.7225
8.1 | -102.724
8.1036 | -102.7321
8.1073 | -102.7423
8.111 | -102.7522
8.1147 | -102.7661
8.1184 | -102.7815
8.1221 | -102.789
8.1258 | -102.8008
8.1295 | -102.8151
8.1332 | -102.8286
8.1368 | -102.8475
8.1405 | -102.8681
8.1442 | -102.8895
8.1479 | -102.9068
8.1516 | -102.9277
8.1553 | -102.9545
8.159 | -102.9751
8.1627 | -103.0004
8.1664 | -103.0269
8.17 | -103.0506
8.1737 | -103.0762
8.1774 | -103.0992
8.1811 | -103.1208
8.1848 | -103.1493
8.1885 | -103.1777
8.1922 | -103.2006
8.1959 | -103.2221
8.1995 | -103.2388
8.2032 | -103.252
8.2069 | -103.2611
8.2106 | -103.2681
8.2143 | -103.2648
8.218 | -103.2585
8.2217 | -103.2465
8.2254 | -103.2273
8.2291 | -103.2057
8.2327 | -103.1803
8.2364 | -103.1486
8.2401 | -103.1173
8.2438 | -103.0865
8.2475 | -103.0567
8.2512 | -103.0319
8.2549 | -103.0107
8.2586 | -102.9911
8.2623 | -102.9691
8.2659 | -102.9522
8.2696 | -102.9473
8.2733 | -102.9444
8.277 | -102.9414
8.2807 | -102.9397
8.2844 | -102.9365
8.2881 | -102.9324
8.2918 | -102.9262
8.2954 | -102.9198
8.2991 | -102.9165
8.3028 | -102.907
8.3065 | -102.8927
8.3102 | -102.876
8.3139 | -102.8601
8.3176 | -102.8441
8.3213 | -102.828
8.325 | -102.8103
8.3286 | -102.79
8.3323 | -102.775
8.336 | -102.7663
8.3397 | -102.7553
8.3434 | -102.7508
8.3471 | -102.7488
8.3508 | -102.7459
8.3545 | -102.7456
8.3582 | -102.7503
8.3618 | -102.7513
8.3655 | -102.759
8.3692 | -102.7685
8.3729 | -102.7795
8.3766 | -102.794
8.3803 | -102.8168
8.384 | -102.8476
8.3877 | -102.881
8.3913 | -102.9125
8.395 | -102.9528
8.3987 | -102.9977
8.4024 | -103.0466
8.4061 | -103.0973
8.4098 | -103.144
8.4135 | -103.1879
8.4172 | -103.2242
8.4209 | -103.2523
8.4245 | -103.276
8.4282 | -103.2918
8.4319 | -103.302
8.4356 | -103.3036
8.4393 | -103.2974
8.443 | -103.2875
8.4467 | -103.2749
8.4504 | -103.2591
8.4541 | -103.2318
8.4577 | -103.2028
8.4614 | -103.1758
8.4651 | -103.1458
8.4688 | -103.1157
8.4725 | -103.0924
8.4762 | -103.0767
8.4799 | -103.061
8.4836 | -103.0509
8.4873 | -103.053
8.4909 | -103.0685
8.4946 | -103.092
8.4983 | -103.1237
8.502 | -103.1649
8.5057 | -103.2179
8.5094 | -103.2712
8.5131 | -103.3239
8.5168 | -103.3756
8.5204 | -103.4251
8.5241 | -103.4714
8.5278 | -103.5127
8.5315 | -103.543
8.5352 | -103.5681
8.5389 | -103.5826
8.5426 | -103.591
8.5463 | -103.5932
8.55 | -103.5872
8.5536 | -103.573
8.5573 | -103.5525
8.561 | -103.5232
8.5647 | -103.4806
8.5684 | -103.4337
8.5721 | -103.3879
8.5758 | -103.3424
8.5795 | -103.3003
8.5832 | -103.2541
8.5868 | -103.208
8.5905 | -103.1624
8.5942 | -103.1215
8.5979 | -103.0862
8.6016 | -103.0546
8.6053 | -103.033
8.609 | -103.0178
8.6127 | -102.9959
8.6163 | -102.9746
8.62 | -102.9586
8.6237 | -102.9507
8.6274 | -102.9453
8.6311 | -102.9428
8.6348 | -102.9508
8.6385 | -102.9655
8.6422 | -102.9848
8.6459 | -103.0107
8.6495 | -103.042
8.6532 | -103.0765
8.6569 | -103.1173
8.6606 | -103.1687
8.6643 | -103.2243
8.668 | -103.2821
8.6717 | -103.3351
8.6754 | -103.3851
8.6791 | -103.4338
8.6827 | -103.4819
8.6864 | -103.5252
8.6901 | -103.5627
8.6938 | -103.5945
8.6975 | -103.6118
8.7012 | -103.6133
8.7049 | -103.6063
8.7086 | -103.5931
8.7122 | -103.5774
8.7159 | -103.5563
8.7196 | -103.5345
8.7233 | -103.5078
8.727 | -103.4692
8.7307 | -103.4366
8.7344 | -103.4104
8.7381 | -103.3893
8.7418 | -103.3776
8.7454 | -103.3688
8.7491 | -103.3581
8.7528 | -103.3508
8.7565 | -103.3464
8.7602 | -103.3441
8.7639 | -103.3463
8.7676 | -103.3488
8.7713 | -103.3495
8.775 | -103.3531
8.7786 | -103.3581
8.7823 | -103.3705
8.786 | -103.3799
8.7897 | -103.3867
8.7934 | -103.3929
8.7971 | -103.4017
8.8008 | -103.4139
8.8045 | -103.4321
8.8082 | -103.4495
8.8118 | -103.4625
8.8155 | -103.479
8.8192 | -103.5005
8.8229 | -103.5186
8.8266 | -103.537
8.8303 | -103.557
8.834 | -103.5771
8.8377 | -103.5912
8.8413 | -103.5963
8.845 | -103.5985
8.8487 | -103.5964
8.8524 | -103.591
8.8561 | -103.5899
8.8598 | -103.5843
8.8635 | -103.5753
8.8672 | -103.5609
8.8709 | -103.5495
8.8745 | -103.5426
8.8782 | -103.5401
8.8819 | -103.5349
8.8856 | -103.5306
8.8893 | -103.5249
8.893 | -103.5136
8.8967 | -103.5003
8.9004 | -103.4837
8.9041 | -103.4635
8.9077 | -103.4411
8.9114 | -103.4135
8.9151 | -103.3816
8.9188 | -103.3453
8.9225 | -103.3064
8.9262 | -103.2665
8.9299 | -103.2285
8.9336 | -103.189
8.9372 | -103.1547
8.9409 | -103.1252
8.9446 | -103.0981
8.9483 | -103.0724
8.952 | -103.0581
8.9557 | -103.0491
8.9594 | -103.043
8.9631 | -103.038
8.9668 | -103.0403
8.9704 | -103.0515
8.9741 | -103.0666
8.9778 | -103.0886
8.9815 | -103.1141
8.9852 | -103.1447
8.9889 | -103.1797
8.9926 | -103.2176
8.9963 | -103.2648
9 | -103.3095
9.0036 | -103.3551
9.0073 | -103.4022
9.011 | -103.4449
9.0147 | -103.4848
9.0184 | -103.5196
9.0221 | -103.5506
9.0258 | -103.5821
9.0295 | -103.6058
9.0331 | -103.6179
9.0368 | -103.621
9.0405 | -103.6194
9.0442 | -103.6066
9.0479 | -103.6041
9.0516 | -103.5978
9.0553 | -103.591
9.059 | -103.582
9.0627 | -103.5721
9.0663 | -103.565
9.07 | -103.56
9.0737 | -103.5595
9.0774 | -103.562
9.0811 | -103.5713
9.0848 | -103.585
9.0885 | -103.5997
9.0922 | -103.6203
9.0959 | -103.6469
9.0995 | -103.6772
9.1032 | -103.7072
9.1069 | -103.7354
9.1106 | -103.7607
9.1143 | -103.7859
9.118 | -103.8103
9.1217 | -103.83
9.1254 | -103.8461
9.1291 | -103.8605
9.1327 | -103.8755
9.1364 | -103.8868
9.1401 | -103.8921
9.1438 | -103.8983
9.1475 | -103.9045
9.1512 | -103.9094
9.1549 | -103.9095
9.1586 | -103.9007
9.1622 | -103.8911
9.1659 | -103.8871
9.1696 | -103.8793
9.1733 | -103.8663
9.177 | -103.8469
9.1807 | -103.8275
9.1844 | -103.8108
9.1881 | -103.7953
9.1918 | -103.7752
9.1954 | -103.7548
9.1991 | -103.7323
9.2028 | -103.7117
9.2065 | -103.6965
9.2102 | -103.6861
9.2139 | -103.6792
9.2176 | -103.6752
9.2213 | -103.6734
9.225 | -103.6817
9.2286 | -103.6968
9.2323 | -103.7175
9.236 | -103.7426
9.2397 | -103.7722
9.2434 | -103.8064
9.2471 | -103.8424
9.2508 | -103.8818
9.2545 | -103.9172
9.2581 | -103.9552
9.2618 | -103.9961
9.2655 | -104.0321
9.2692 | -104.0647
9.2729 | -104.0921
9.2766 | -104.1131
9.2803 | -104.1249
9.284 | -104.128
9.2877 | -104.1239
9.2913 | -104.1114
9.295 | -104.0939
9.2987 | -104.0696
9.3024 | -104.032
9.3061 | -103.9908
9.3098 | -103.9467
9.3135 | -103.8996
9.3172 | -103.8509
9.3209 | -103.7974
9.3245 | -103.7437
9.3282 | -103.6961
9.3319 | -103.6535
9.3356 | -103.616
9.3393 | -103.5844
9.343 | -103.554
9.3467 | -103.5253
9.3504 | -103.5024
9.3541 | -103.4858
9.3577 | -103.4722
9.3614 | -103.459
9.3651 | -103.4457
9.3688 | -103.4335
9.3725 | -103.4275
9.3762 | -103.4265
9.3799 | -103.4263
9.3836 | -103.4278
9.3872 | -103.428
9.3909 | -103.4323
9.3946 | -103.4346
9.3983 | -103.4377
9.402 | -103.4472
9.4057 | -103.4588
9.4094 | -103.4761
9.4131 | -103.4978
9.4168 | -103.519
9.4204 | -103.5421
9.4241 | -103.5676
9.4278 | -103.5953
9.4315 | -103.6215
9.4352 | -103.6479
9.4389 | -103.6786
9.4426 | -103.7103
9.4463 | -103.7421
9.45 | -103.7729
9.4536 | -103.7998
9.4573 | -103.8304
9.461 | -103.8581
9.4647 | -103.8836
9.4684 | -103.9081
9.4721 | -103.932
9.4758 | -103.955
9.4795 | -103.9796
9.4831 | -104.004
9.4868 | -104.0258
9.4905 | -104.0439
9.4942 | -104.0642
9.4979 | -104.0856
9.5016 | -104.1051
9.5053 | -104.1217
9.509 | -104.1289
9.5127 | -104.1346
9.5163 | -104.1397
9.52 | -104.1415
9.5237 | -104.1401
9.5274 | -104.1293
9.5311 | -104.1133
9.5348 | -104.0931
9.5385 | -104.0688
9.5422 | -104.042
9.5459 | -104.0133
9.5495 | -103.9885
9.5532 | -103.9697
9.5569 | -103.9514
9.5606 | -103.9353
9.5643 | -103.9268
9.568 | -103.9292
9.5717 | -103.9401
9.5754 | -103.9759
9.579 | -103.9883
9.5827 | -104.0052
9.5864 | -104.0246
9.5901 | -104.049
9.5938 | -104.0775
9.5975 | -104.1053
9.6012 | -104.1281
9.6049 | -104.1446
9.6086 | -104.1552
9.6122 | -104.1622
9.6159 | -104.1632
9.6196 | -104.1644
9.6233 | -104.1601
9.627 | -104.1471
9.6307 | -104.1312
9.6344 | -104.114
9.6381 | -104.0926
9.6418 | -104.0682
9.6454 | -104.044
9.6491 | -104.0204
9.6528 | -103.9896
9.6565 | -103.9603
9.6602 | -103.9326
9.6639 | -103.9039
9.6676 | -103.8755
9.6713 | -103.8522
9.675 | -103.8299
9.6786 | -103.8084
9.6823 | -103.7909
9.686 | -103.7831
9.6897 | -103.7789
9.6934 | -103.7769
9.6971 | -103.7794
9.7008 | -103.7886
9.7045 | -103.7992
9.7081 | -103.8113
9.7118 | -103.8246
9.7155 | -103.8433
9.7192 | -103.867
9.7229 | -103.8902
9.7266 | -103.9107
9.7303 | -103.9319
9.734 | -103.9472
9.7377 | -103.9573
9.7413 | -103.9681
9.745 | -103.9836
9.7487 | -103.9942
9.7524 | -104.001
9.7561 | -104.0076
9.7598 | -104.0135
9.7635 | -104.0143
9.7672 | -104.0172
9.7709 | -104.0217
9.7745 | -104.0236
9.7782 | -104.0253
9.7819 | -104.0299
9.7856 | -104.0411
9.7893 | -104.0538
9.793 | -104.068
9.7967 | -104.0852
9.8004 | -104.1008
9.804 | -104.1149
9.8077 | -104.1332
9.8114 | -104.1582
9.8151 | -104.1794
9.8188 | -104.2016
9.8225 | -104.2234
9.8262 | -104.2423
9.8299 | -104.2559
9.8336 | -104.265
9.8372 | -104.273
9.8409 | -104.2818
9.8446 | -104.2879
9.8483 | -104.2939
9.852 | -104.294
9.8557 | -104.2929
9.8594 | -104.2898
9.8631 | -104.2873
9.8668 | -104.2888
9.8704 | -104.2889
9.8741 | -104.285
9.8778 | -104.2872
9.8815 | -104.2949
9.8852 | -104.3008
9.8889 | -104.3096
9.8926 | -104.3264
9.8963 | -104.345
9.8999 | -104.3632
9.9036 | -104.383
9.9073 | -104.4068
9.911 | -104.4339
9.9147 | -104.4611
9.9184 | -104.4892
9.9221 | -104.5081
9.9258 | -104.5245
9.9295 | -104.5402
9.9331 | -104.551
9.9368 | -104.557
9.9405 | -104.5625
9.9442 | -104.5611
9.9479 | -104.554
9.9516 | -104.5425
9.9553 | -104.5248
9.959 | -104.5045
9.9627 | -104.4802
9.9663 | -104.4545
9.97 | -104.4299
9.9737 | -104.4081
9.9774 | -104.3932
9.9811 | -104.3815
9.9848 | -104.3695
9.9885 | -104.3516
9.9922 | -104.3375
9.9959 | -104.3288
9.9995 | -104.3215
10.0032 | -104.3127
10.0069 | -104.3042
10.0106 | -104.29
10.0143 | -104.2745
10.018 | -104.257
10.0217 | -104.2378
10.0254 | -104.2175
10.029 | -104.1989
10.0327 | -104.1769
10.0364 | -104.1577
10.0401 | -104.1423
10.0438 | -104.1288
10.0475 | -104.1175
10.0512 | -104.1101
10.0549 | -104.1041
10.0586 | -104.1081
10.0622 | -104.1172
10.0659 | -104.1294
10.0696 | -104.1438
10.0733 | -104.1582
10.077 | -104.177
10.0807 | -104.1972
10.0844 | -104.2123
10.0881 | -104.2302
10.0918 | -104.2463
10.0954 | -104.2545
10.0991 | -104.2577
10.1028 | -104.2557
10.1065 | -104.2451
10.1102 | -104.2091
10.1139 | -104.1949
10.1176 | -104.172
10.1213 | -104.1441
10.1249 | -104.1157
10.1286 | -104.0866
10.1323 | -104.059
10.136 | -104.0352
10.1397 | -104.0079
10.1434 | -103.9866
10.1471 | -103.9694
10.1508 | -103.9559
10.1545 | -103.9468
10.1581 | -103.9443
10.1618 | -103.9458
10.1655 | -103.9527
10.1692 | -103.961
10.1729 | -103.9711
10.1766 | -103.9849
10.1803 | -104.0038
10.184 | -104.0245
10.1877 | -104.0479
10.1913 | -104.0725
10.195 | -104.1009
10.1987 | -104.1288
10.2024 | -104.1531
10.2061 | -104.1747
10.2098 | -104.1962
10.2135 | -104.2239
10.2172 | -104.2508
10.2208 | -104.2719
10.2245 | -104.2893
10.2282 | -104.3085
10.2319 | -104.3255
10.2356 | -104.335
10.2393 | -104.3401
10.243 | -104.3446
10.2467 | -104.3471
10.2504 | -104.3463
10.254 | -104.3476
10.2577 | -104.3455
10.2614 | -104.3403
10.2651 | -104.3327
10.2688 | -104.3261
10.2725 | -104.3203
10.2762 | -104.3116
10.2799 | -104.3009
10.2836 | -104.2908
10.2872 | -104.2842
10.2909 | -104.2784
10.2946 | -104.2725
10.2983 | -104.2677
10.302 | -104.2631
10.3057 | -104.2608
10.3094 | -104.2613
10.3131 | -104.2637
10.3168 | -104.2661
10.3204 | -104.2663
10.3241 | -104.2635
10.3278 | -104.257
10.3315 | -104.2584
10.3352 | -104.2607
10.3389 | -104.2629
10.3426 | -104.2665
10.3463 | -104.266
10.3499 | -104.2654
10.3536 | -104.2728
10.3573 | -104.2744
10.361 | -104.2753
10.3647 | -104.2748
10.3684 | -104.2749
10.3721 | -104.2818
10.3758 | -104.2871
10.3795 | -104.2909
10.3831 | -104.2947
10.3868 | -104.3026
10.3905 | -104.3106
10.3942 | -104.3188
10.3979 | -104.328
10.4016 | -104.3401
10.4053 | -104.3516
10.409 | -104.3635
10.4127 | -104.376
10.4163 | -104.3901
10.42 | -104.4105
10.4237 | -104.429
10.4274 | -104.4476
10.4311 | -104.4666
10.4348 | -104.4843
10.4385 | -104.5021
10.4422 | -104.5185
10.4458 | -104.5302
10.4495 | -104.5387
10.4532 | -104.5516
10.4569 | -104.5613
10.4606 | -104.5674
10.4643 | -104.572
10.468 | -104.5712
10.4717 | -104.5709
10.4754 | -104.5735
10.479 | -104.5737
10.4827 | -104.5576
10.4864 | -104.5407
10.4901 | -104.5218
10.4938 | -104.503
10.4975 | -104.4846
10.5012 | -104.4626
10.5049 | -104.4409
10.5086 | -104.4265
10.5122 | -104.4102
10.5159 | -104.3959
10.5196 | -104.3814
10.5233 | -104.3666
10.527 | -104.3606
10.5307 | -104.3571
10.5344 | -104.3604
10.5381 | -104.3593
10.5417 | -104.3602
10.5454 | -104.3598
10.5491 | -104.3568
10.5528 | -104.3564
10.5565 | -104.3541
10.5602 | -104.3479
10.5639 | -104.3405
10.5676 | -104.3312
10.5713 | -104.3216
10.5749 | -104.3109
10.5786 | -104.3023
10.5823 | -104.2967
10.586 | -104.2936
10.5897 | -104.2895
10.5934 | -104.2859
10.5971 | -104.2816
10.6008 | -104.2786
10.6045 | -104.2731
10.6081 | -104.2713
10.6118 | -104.2689
10.6155 | -104.267
10.6192 | -104.2671
10.6229 | -104.2627
10.6266 | -104.2572
10.6303 | -104.2532
10.634 | -104.2477
10.6377 | -104.239
10.6413 | -104.2321
10.645 | -104.226
10.6487 | -104.2215
10.6524 | -104.2185
10.6561 | -104.2177
10.6598 | -104.2157
10.6635 | -104.211
10.6672 | -104.2097
10.6708 | -104.211
10.6745 | -104.2161
10.6782 | -104.2241
10.6819 | -104.2329
10.6856 | -104.2411
10.6893 | -104.2482
10.693 | -104.2646
10.6967 | -104.2805
10.7004 | -104.3012
10.704 | -104.3272
10.7077 | -104.3528
10.7114 | -104.3788
10.7151 | -104.4062
10.7188 | -104.4352
10.7225 | -104.4656
10.7262 | -104.4926
10.7299 | -104.5207
10.7336 | -104.5468
10.7372 | -104.572
10.7409 | -104.5946
10.7446 | -104.6159
10.7483 | -104.6357
10.752 | -104.6542
10.7557 | -104.6713
10.7594 | -104.6823
10.7631 | -104.6856
10.7667 | -104.6864
10.7704 | -104.686
10.7741 | -104.6838
10.7778 | -104.6827
10.7815 | -104.6786
10.7852 | -104.6731
10.7889 | -104.6678
10.7926 | -104.6605
10.7963 | -104.6503
10.7999 | -104.642
10.8036 | -104.637
10.8073 | -104.6319
10.811 | -104.6273
10.8147 | -104.6233
10.8184 | -104.6205
10.8221 | -104.62
10.8258 | -104.6215
10.8295 | -104.6231
10.8331 | -104.6255
10.8368 | -104.6292
10.8405 | -104.6319
10.8442 | -104.6345
10.8479 | -104.6374
10.8516 | -104.6363
10.8553 | -104.636
10.859 | -104.6346
10.8626 | -104.6339
10.8663 | -104.6311
10.87 | -104.6297
10.8737 | -104.6309
10.8774 | -104.6282
10.8811 | -104.6222
10.8848 | -104.6171
10.8885 | -104.6122
10.8922 | -104.6078
10.8958 | -104.5986
10.8995 | -104.5858
10.9032 | -104.5763
10.9069 | -104.5669
10.9106 | -104.5547
10.9143 | -104.542
10.918 | -104.5325
10.9217 | -104.5242
10.9254 | -104.5191
10.929 | -104.5146
10.9327 | -104.505
10.9364 | -104.4976
10.9401 | -104.4937
10.9438 | -104.4936
10.9475 | -104.4971
10.9512 | -104.4997
10.9549 | -104.5036
10.9586 | -104.5054
10.9622 | -104.5079
10.9659 | -104.5152
10.9696 | -104.5266
10.9733 | -104.5393
10.977 | -104.5519
10.9807 | -104.5655
10.9844 | -104.5785
10.9881 | -104.5904
10.9917 | -104.6013
10.9954 | -104.6118
10.9991 | -104.6272
11.0028 | -104.6457
11.0065 | -104.6614
11.0102 | -104.6724
11.0139 | -104.6825
11.0176 | -104.687
11.0213 | -104.6843
11.0249 | -104.6837
11.0286 | -104.6837
11.0323 | -104.6795
11.036 | -104.6744
11.0397 | -104.6702
11.0434 | -104.663
11.0471 | -104.6545
11.0508 | -104.6436
11.0545 | -104.6344
11.0581 | -104.6282
11.0618 | -104.6181
11.0655 | -104.6094
11.0692 | -104.6003
11.0729 | -104.5897
11.0766 | -104.5808
11.0803 | -104.5704
11.084 | -104.5613
11.0876 | -104.5516
11.0913 | -104.5473
11.095 | -104.542
11.0987 | -104.5373
11.1024 | -104.5333
11.1061 | -104.534
11.1098 | -104.5389
11.1135 | -104.5452
11.1172 | -104.5521
11.1208 | -104.5602
11.1245 | -104.5683
11.1282 | -104.5763
11.1319 | -104.5886
11.1356 | -104.6059
11.1393 | -104.6244
11.143 | -104.6422
11.1467 | -104.659
11.1504 | -104.6778
11.154 | -104.695
11.1577 | -104.7104
11.1614 | -104.7239
11.1651 | -104.7385
11.1688 | -104.7547
11.1725 | -104.7626
11.1762 | -104.7741
11.1799 | -104.7854
11.1835 | -104.7989
11.1872 | -104.8066
11.1909 | -104.8096
11.1946 | -104.8119
11.1983 | -104.8172
11.202 | -104.8153
11.2057 | -104.8121
11.2094 | -104.8101
11.2131 | -104.8081
11.2167 | -104.8091
11.2204 | -104.8075
11.2241 | -104.8041
11.2278 | -104.8002
11.2315 | -104.7985
11.2352 | -104.7958
11.2389 | -104.7892
11.2426 | -104.7901
11.2463 | -104.7919
11.2499 | -104.7924
11.2536 | -104.7936
11.2573 | -104.7972
11.261 | -104.7997
11.2647 | -104.8047
11.2684 | -104.8109
11.2721 | -104.8185
11.2758 | -104.8221
11.2795 | -104.8258
11.2831 | -104.8312
11.2868 | -104.8317
11.2905 | -104.8309
11.2942 | -104.8301
11.2979 | -104.8234
11.3016 | -104.8169
11.3053 | -104.808
11.309 | -104.7975
11.3126 | -104.7855
11.3163 | -104.7717
11.32 | -104.752
11.3237 | -104.7322
11.3274 | -104.71
11.3311 | -104.6855
11.3348 | -104.6625
11.3385 | -104.6395
11.3422 | -104.6158
11.3458 | -104.5913
11.3495 | -104.5657
11.3532 | -104.545
11.3569 | -104.5281
11.3606 | -104.5108
11.3643 | -104.4954
11.368 | -104.4847
11.3717 | -104.4786
11.3754 | -104.4732
11.379 | -104.4677
11.3827 | -104.4605
11.3864 | -104.4584
11.3901 | -104.4586
11.3938 | -104.463
11.3975 | -104.4735
11.4012 | -104.4835
11.4049 | -104.4941
11.4085 | -104.5034
11.4122 | -104.5114
11.4159 | -104.5212
11.4196 | -104.53
11.4233 | -104.5353
11.427 | -104.5398
11.4307 | -104.5426
11.4344 | -104.5431
11.4381 | -104.5426
11.4417 | -104.5426
11.4454 | -104.5419
11.4491 | -104.5374
11.4528 | -104.5317
11.4565 | -104.5225
11.4602 | -104.5043
11.4639 | -104.4899
11.4676 | -104.4774
11.4713 | -104.4624
11.4749 | -104.4456
11.4786 | -104.427
11.4823 | -104.4084
11.486 | -104.3936
11.4897 | -104.3786
11.4934 | -104.3627
11.4971 | -104.3509
11.5008 | -104.3378
11.5044 | -104.3289
11.5081 | -104.3242
11.5118 | -104.3205
11.5155 | -104.3205
11.5192 | -104.3239
11.5229 | -104.3289
11.5266 | -104.3393
11.5303 | -104.3505
11.534 | -104.3667
11.5376 | -104.3873
11.5413 | -104.409
11.545 | -104.4363
11.5487 | -104.4652
11.5524 | -104.4935
11.5561 | -104.5252
11.5598 | -104.555
11.5635 | -104.5795
11.5672 | -104.6079
11.5708 | -104.6353
11.5745 | -104.6581
11.5782 | -104.679
11.5819 | -104.6986
11.5856 | -104.7163
11.5893 | -104.7328
11.593 | -104.7448
11.5967 | -104.7478
11.6004 | -104.7484
11.604 | -104.7445
11.6077 | -104.737
11.6114 | -104.7229
11.6151 | -104.7041
11.6188 | -104.6842
11.6225 | -104.6637
11.6262 | -104.6457
11.6299 | -104.6211
11.6335 | -104.5923
11.6372 | -104.5679
11.6409 | -104.5446
11.6446 | -104.5211
11.6483 | -104.4982
11.652 | -104.4732
11.6557 | -104.4535
11.6594 | -104.4365
11.6631 | -104.4186
11.6667 | -104.4047
11.6704 | -104.3933
11.6741 | -104.3831
11.6778 | -104.3777
11.6815 | -104.3752
11.6852 | -104.3723
11.6889 | -104.3742
11.6926 | -104.3778
11.6963 | -104.3846
11.6999 | -104.3958
11.7036 | -104.4107
11.7073 | -104.4212
11.711 | -104.4334
11.7147 | -104.4459
11.7184 | -104.4614
11.7221 | -104.4775
11.7258 | -104.4943
11.7294 | -104.5107
11.7331 | -104.5281
11.7368 | -104.5488
11.7405 | -104.5711
11.7442 | -104.5912
11.7479 | -104.6079
11.7516 | -104.6231
11.7553 | -104.6408
11.759 | -104.6541
11.7626 | -104.6698
11.7663 | -104.688
11.77 | -104.7046
11.7737 | -104.721
11.7774 | -104.7351
11.7811 | -104.7465
11.7848 | -104.7558
11.7885 | -104.7634
11.7922 | -104.7699
11.7958 | -104.773
11.7995 | -104.7705
11.8032 | -104.769
11.8069 | -104.7668
11.8106 | -104.7612
11.8143 | -104.7532
11.818 | -104.7458
11.8217 | -104.7396
11.8253 | -104.7257
11.829 | -104.7131
11.8327 | -104.7026
11.8364 | -104.6912
11.8401 | -104.678
11.8438 | -104.6681
11.8475 | -104.6582
11.8512 | -104.6476
11.8549 | -104.6409
11.8585 | -104.6348
11.8622 | -104.6294
11.8659 | -104.627
11.8696 | -104.6251
11.8733 | -104.6279
11.877 | -104.6295
11.8807 | -104.6319
11.8844 | -104.6354
11.8881 | -104.6373
11.8917 | -104.6405
11.8954 | -104.645
11.8991 | -104.6506
11.9028 | -104.6594
11.9065 | -104.6719
11.9102 | -104.6796
11.9139 | -104.6861
11.9176 | -104.6966
11.9213 | -104.7078
11.9249 | -104.7179
11.9286 | -104.7227
11.9323 | -104.7279
11.936 | -104.7297
11.9397 | -104.7297
11.9434 | -104.7294
11.9471 | -104.7267
11.9508 | -104.7258
11.9544 | -104.7249
11.9581 | -104.7224
11.9618 | -104.719
11.9655 | -104.7141
11.9692 | -104.708
11.9729 | -104.7035
11.9766 | -104.7002
11.9803 | -104.695
11.984 | -104.6904
11.9876 | -104.6878
11.9913 | -104.6836
11.995 | -104.6781
11.9987 | -104.6755
12.0024 | -104.6722
12.0061 | -104.6681
12.0098 | -104.6645
12.0135 | -104.6594
12.0172 | -104.655
12.0208 | -104.6519
12.0245 | -104.6501
12.0282 | -104.6459
12.0319 | -104.645
12.0356 | -104.6452
12.0393 | -104.6452
12.043 | -104.6459
12.0467 | -104.6481
12.0503 | -104.6547
12.054 | -104.6616
12.0577 | -104.6708
12.0614 | -104.6816
12.0651 | -104.6908
12.0688 | -104.6992
12.0725 | -104.7079
12.0762 | -104.7127
12.0799 | -104.7117
12.0835 | -104.714
12.0872 | -104.7201
12.0909 | -104.7218
12.0946 | -104.7207
12.0983 | -104.7176
12.102 | -104.7145
12.1057 | -104.7078
12.1094 | -104.7015
12.1131 | -104.6948
12.1167 | -104.6848
12.1204 | -104.6724
12.1241 | -104.6605
12.1278 | -104.6482
12.1315 | -104.6369
12.1352 | -104.6249
12.1389 | -104.618
12.1426 | -104.6156
12.1462 | -104.6161
12.1499 | -104.6135
12.1536 | -104.6107
12.1573 | -104.612
12.161 | -104.6158
12.1647 | -104.6223
12.1684 | -104.6327
12.1721 | -104.6426
12.1758 | -104.6549
12.1794 | -104.6682
12.1831 | -104.6863
12.1868 | -104.706
12.1905 | -104.7251
12.1942 | -104.7437
12.1979 | -104.7607
12.2016 | -104.7797
12.2053 | -104.7945
12.209 | -104.8092
12.2126 | -104.8262
12.2163 | -104.8407
12.22 | -104.8544
12.2237 | -104.8641
12.2274 | -104.8719
12.2311 | -104.8804
12.2348 | -104.8879
12.2385 | -104.8902
12.2422 | -104.8978
12.2458 | -104.907
12.2495 | -104.9127
12.2532 | -104.9215
12.2569 | -104.93
12.2606 | -104.9358
12.2643 | -104.9406
12.268 | -104.9456
12.2717 | -104.9542
12.2753 | -104.96
12.279 | -104.9664
12.2827 | -104.9674
12.2864 | -104.9656
12.2901 | -104.9666
12.2938 | -104.9678
12.2975 | -104.9684
12.3012 | -104.9662
12.3049 | -104.9627
12.3085 | -104.9564
12.3122 | -104.9484
12.3159 | -104.9379
12.3196 | -104.9236
12.3233 | -104.9094
12.327 | -104.8957
12.3307 | -104.8831
12.3344 | -104.8696
12.3381 | -104.8524
12.3417 | -104.834
12.3454 | -104.8168
12.3491 | -104.8034
12.3528 | -104.7887
12.3565 | -104.7772
12.3602 | -104.7673
12.3639 | -104.7612
12.3676 | -104.7591
12.3712 | -104.7579
12.3749 | -104.7576
12.3786 | -104.7619
12.3823 | -104.7706
12.386 | -104.7805
12.3897 | -104.7917
12.3934 | -104.8041
12.3971 | -104.8173
12.4008 | -104.8355
12.4044 | -104.8532
12.4081 | -104.8698
12.4118 | -104.8868
12.4155 | -104.9042
12.4192 | -104.9155
12.4229 | -104.9258
12.4266 | -104.94
12.4303 | -104.9561
12.434 | -104.9705
12.4376 | -104.9865
12.4413 | -104.9988
12.445 | -105.0073
12.4487 | -105.0155
12.4524 | -105.0222
12.4561 | -105.0303
12.4598 | -105.0363
12.4635 | -105.0469
12.4671 | -105.0544
12.4708 | -105.0558
12.4745 | -105.0582
12.4782 | -105.0631
12.4819 | -105.0657
12.4856 | -105.0683
12.4893 | -105.0722
12.493 | -105.0747
12.4967 | -105.0773
12.5003 | -105.0832
12.504 | -105.0893
12.5077 | -105.0946
12.5114 | -105.0975
12.5151 | -105.0988
12.5188 | -105.1061
12.5225 | -105.1131
12.5262 | -105.1194
12.5299 | -105.124
12.5335 | -105.125
12.5372 | -105.1223
12.5409 | -105.1193
12.5446 | -105.1156
12.5483 | -105.1112
12.552 | -105.1103
12.5557 | -105.1114
12.5594 | -105.1092
12.5631 | -105.1058
12.5667 | -105.1015
12.5704 | -105.0973
12.5741 | -105.0967
12.5778 | -105.0971
12.5815 | -105.0967
12.5852 | -105.0975
12.5889 | -105.0973
12.5926 | -105.0989
12.5962 | -105.103
12.5999 | -105.1099
12.6036 | -105.1149
12.6073 | -105.1146
12.611 | -105.1152
12.6147 | -105.1185
12.6184 | -105.1244
12.6221 | -105.1301
12.6258 | -105.1317
12.6294 | -105.1376
12.6331 | -105.1438
12.6368 | -105.1495
12.6405 | -105.1535
12.6442 | -105.1606
12.6479 | -105.1662
12.6516 | -105.1723
12.6553 | -105.1787
12.659 | -105.1805
12.6626 | -105.1839
12.6663 | -105.1879
12.67 | -105.1903
12.6737 | -105.1939
12.6774 | -105.1951
12.6811 | -105.1939
12.6848 | -105.1926
12.6885 | -105.1928
12.6921 | -105.1951
12.6958 | -105.1967
12.6995 | -105.1916
12.7032 | -105.1868
12.7069 | -105.1798
12.7106 | -105.1729
12.7143 | -105.1714
12.718 | -105.1706
12.7217 | -105.1678
12.7253 | -105.1656
12.729 | -105.1606
12.7327 | -105.1525
12.7364 | -105.1468
12.7401 | -105.1436
12.7438 | -105.1418
12.7475 | -105.1411
12.7512 | -105.14
12.7549 | -105.1402
12.7585 | -105.1389
12.7622 | -105.1374
12.7659 | -105.1362
12.7696 | -105.1364
12.7733 | -105.1384
12.777 | -105.1415
12.7807 | -105.1434
12.7844 | -105.144
12.788 | -105.1462
12.7917 | -105.1477
12.7954 | -105.1482
12.7991 | -105.1466
12.8028 | -105.1451
12.8065 | -105.1469
12.8102 | -105.1533
12.8139 | -105.1518
12.8176 | -105.1492
12.8212 | -105.1495
12.8249 | -105.1528
12.8286 | -105.1539
12.8323 | -105.1574
12.836 | -105.1594
12.8397 | -105.1592
12.8434 | -105.1617
12.8471 | -105.1672
12.8508 | -105.1711
12.8544 | -105.1754
12.8581 | -105.1796
12.8618 | -105.1863
12.8655 | -105.1949
12.8692 | -105.2037
12.8729 | -105.2095
12.8766 | -105.2141
12.8803 | -105.2192
12.884 | -105.2283
12.8876 | -105.2362
12.8913 | -105.2447
12.895 | -105.2537
12.8987 | -105.259
12.9024 | -105.2658
12.9061 | -105.2717
12.9098 | -105.2779
12.9135 | -105.2859
12.9171 | -105.2956
12.9208 | -105.2998
12.9245 | -105.2964
12.9282 | -105.2933
12.9319 | -105.2899
12.9356 | -105.286
12.9393 | -105.2787
12.943 | -105.2718
12.9467 | -105.2636
12.9503 | -105.2551
12.954 | -105.2434
12.9577 | -105.2295
12.9614 | -105.2093
12.9651 | -105.1857
12.9688 | -105.1636
12.9725 | -105.1403
12.9762 | -105.1144
12.9799 | -105.0878
12.9835 | -105.0615
12.9872 | -105.0349
12.9909 | -105.0047
12.9946 | -104.9774
12.9983 | -104.9499
13.002 | -104.9269
13.0057 | -104.9062
13.0094 | -104.8848
13.013 | -104.8654
13.0167 | -104.8513
13.0204 | -104.8412
13.0241 | -104.8343
13.0278 | -104.8308
13.0315 | -104.8282
13.0352 | -104.8239
13.0389 | -104.8224
13.0426 | -104.8252
13.0462 | -104.8291
13.0499 | -104.8335
13.0536 | -104.8385
13.0573 | -104.8435
13.061 | -104.8539
13.0647 | -104.8674
13.0684 | -104.8813
13.0721 | -104.8941
13.0758 | -104.9074
13.0794 | -104.9206
13.0831 | -104.9333
13.0868 | -104.9486
13.0905 | -104.9623
13.0942 | -104.9793
13.0979 | -104.9974
13.1016 | -105.0141
13.1053 | -105.0302
13.1089 | -105.0488
13.1126 | -105.0717
13.1163 | -105.0921
13.12 | -105.1135
13.1237 | -105.1303
13.1274 | -105.1469
13.1311 | -105.1625
13.1348 | -105.1766
13.1385 | -105.1894
13.1421 | -105.199
13.1458 | -105.2087
13.1495 | -105.2215
13.1532 | -105.2317
13.1569 | -105.2368
13.1606 | -105.2411
13.1643 | -105.2439
13.168 | -105.245
13.1717 | -105.2414
13.1753 | -105.236
13.179 | -105.2319
13.1827 | -105.2207
13.1864 | -105.2106
13.1901 | -105.2003
13.1938 | -105.1854
13.1975 | -105.1702
13.2012 | -105.1516
13.2049 | -105.1279
13.2085 | -105.1033
13.2122 | -105.0813
13.2159 | -105.0605
13.2196 | -105.0391
13.2233 | -105.0207
13.227 | -105.0042
13.2307 | -104.9869
13.2344 | -104.9682
13.238 | -104.9506
13.2417 | -104.9375
13.2454 | -104.9277
13.2491 | -104.9167
13.2528 | -104.9066
13.2565 | -104.8995
13.2602 | -104.8941
13.2639 | -104.8924
13.2676 | -104.8981
13.2712 | -104.9043
13.2749 | -104.909
13.2786 | -104.9184
13.2823 | -104.9283
13.286 | -104.9386
13.2897 | -104.9495
13.2934 | -104.9626
13.2971 | -104.9773
13.3008 | -104.994
13.3044 | -105.0085
13.3081 | -105.0228
13.3118 | -105.0356
13.3155 | -105.0504
13.3192 | -105.067
13.3229 | -105.0828
13.3266 | -105.0968
13.3303 | -105.111
13.3339 | -105.1279
13.3376 | -105.1427
13.3413 | -105.1535
13.345 | -105.1646
13.3487 | -105.1747
13.3524 | -105.1846
13.3561 | -105.1914
13.3598 | -105.1977
13.3635 | -105.2057
13.3671 | -105.2132
13.3708 | -105.2145
13.3745 | -105.2159
13.3782 | -105.2147
13.3819 | -105.2084
13.3856 | -105.2035
13.3893 | -105.1981
13.393 | -105.1934
13.3967 | -105.1875
13.4003 | -105.1792
13.404 | -105.1708
13.4077 | -105.1581
13.4114 | -105.146
13.4151 | -105.1313
13.4188 | -105.1165
13.4225 | -105.1033
13.4262 | -105.092
13.4298 | -105.0783
13.4335 | -105.0632
13.4372 | -105.0505
13.4409 | -105.0391
13.4446 | -105.0262
13.4483 | -105.0124
13.452 | -105.0035
13.4557 | -104.9952
13.4594 | -104.9816
13.463 | -104.9734
13.4667 | -104.9674
13.4704 | -104.9581
13.4741 | -104.9524
13.4778 | -104.949
13.4815 | -104.9489
13.4852 | -104.9504
13.4889 | -104.9525
13.4926 | -104.9543
13.4962 | -104.9585
13.4999 | -104.9639
13.5036 | -104.9701
13.5073 | -104.9771
13.511 | -104.9847
13.5147 | -104.9906
13.5184 | -105.0007
13.5221 | -105.0143
13.5258 | -105.0286
13.5294 | -105.0435
13.5331 | -105.0577
13.5368 | -105.0746
13.5405 | -105.0905
13.5442 | -105.1055
13.5479 | -105.1225
13.5516 | -105.138
13.5553 | -105.1519
13.5589 | -105.1683
13.5626 | -105.1819
13.5663 | -105.1981
13.57 | -105.215
13.5737 | -105.2293
13.5774 | -105.2444
13.5811 | -105.257
13.5848 | -105.2688
13.5885 | -105.279
13.5921 | -105.2895
13.5958 | -105.2974
13.5995 | -105.301
13.6032 | -105.3065
13.6069 | -105.314
13.6106 | -105.3184
13.6143 | -105.3218
13.618 | -105.322
13.6217 | -105.3239
13.6253 | -105.3265
13.629 | -105.3248
13.6327 | -105.3202
13.6364 | -105.32
13.6401 | -105.3223
13.6438 | -105.3249
13.6475 | -105.3232
13.6512 | -105.323
13.6548 | -105.3219
13.6585 | -105.3169
13.6622 | -105.3123
13.6659 | -105.3103
13.6696 | -105.3059
13.6733 | -105.3039
13.677 | -105.3028
13.6807 | -105.3014
13.6844 | -105.2976
13.688 | -105.2965
13.6917 | -105.2942
13.6954 | -105.2939
13.6991 | -105.293
13.7028 | -105.2939
13.7065 | -105.2956
13.7102 | -105.2969
13.7139 | -105.2973
13.7176 | -105.3
13.7212 | -105.3013
13.7249 | -105.2985
13.7286 | -105.2995
13.7323 | -105.2997
13.736 | -105.2984
13.7397 | -105.2948
13.7434 | -105.286
13.7471 | -105.2797
13.7507 | -105.2754
13.7544 | -105.266
13.7581 | -105.2569
13.7618 | -105.2461
13.7655 | -105.233
13.7692 | -105.2222
13.7729 | -105.2112
13.7766 | -105.1995
13.7803 | -105.1854
13.7839 | -105.1697
13.7876 | -105.1537
13.7913 | -105.1408
13.795 | -105.1286
13.7987 | -105.1163
13.8024 | -105.1011
13.8061 | -105.0904
13.8098 | -105.0801
13.8135 | -105.0731
13.8171 | -105.0662
13.8208 | -105.0612
13.8245 | -105.0603
13.8282 | -105.0623
13.8319 | -105.0705
13.8356 | -105.0679
13.8393 | -105.0683
13.843 | -105.0694
13.8467 | -105.0735
13.8503 | -105.0764
13.854 | -105.0824
13.8577 | -105.0921
13.8614 | -105.1037
13.8651 | -105.1135
13.8688 | -105.1232
13.8725 | -105.1333
13.8762 | -105.1434
13.8798 | -105.1543
13.8835 | -105.1687
13.8872 | -105.1824
13.8909 | -105.1964
13.8946 | -105.2088
13.8983 | -105.2195
13.902 | -105.2277
13.9057 | -105.2354
13.9094 | -105.245
13.913 | -105.2503
13.9167 | -105.2551
13.9204 | -105.2605
13.9241 | -105.2672
13.9278 | -105.2731
13.9315 | -105.2783
13.9352 | -105.2814
13.9389 | -105.2857
13.9426 | -105.2884
13.9462 | -105.287
13.9499 | -105.2822
13.9536 | -105.2785
13.9573 | -105.2733
13.961 | -105.259
13.9647 | -105.2418
13.9684 | -105.2298
13.9721 | -105.2203
13.9757 | -105.21
13.9794 | -105.2003
13.9831 | -105.1886
13.9868 | -105.1779
13.9905 | -105.1662
13.9942 | -105.1551
13.9979 | -105.1424
14.0016 | -105.1292
14.0053 | -105.1151
14.0089 | -105.1022
14.0126 | -105.0886
14.0163 | -105.0764
14.02 | -105.062
14.0237 | -105.0484
14.0274 | -105.034
14.0311 | -105.0215
14.0348 | -105.0076
14.0385 | -104.9969
14.0421 | -104.9859
14.0458 | -104.9792
14.0495 | -104.9734
14.0532 | -104.9718
14.0569 | -104.972
14.0606 | -104.9724
14.0643 | -104.9718
14.068 | -104.9726
14.0716 | -104.9756
14.0753 | -104.9803
14.079 | -104.9856
14.0827 | -104.994
14.0864 | -105.0053
14.0901 | -105.0185
14.0938 | -105.0315
14.0975 | -105.0432
14.1012 | -105.0602
14.1048 | -105.0799
14.1085 | -105.0981
14.1122 | -105.1149
14.1159 | -105.1303
14.1196 | -105.1461
14.1233 | -105.1598
14.127 | -105.1728
14.1307 | -105.1845
14.1344 | -105.194
14.138 | -105.2054
14.1417 | -105.2167
14.1454 | -105.2271
14.1491 | -105.2369
14.1528 | -105.245
14.1565 | -105.2486
14.1602 | -105.2533
14.1639 | -105.2659
14.1676 | -105.2773
14.1712 | -105.2815
14.1749 | -105.2836
14.1786 | -105.2861
14.1823 | -105.2891
14.186 | -105.2947
14.1897 | -105.2985
14.1934 | -105.3023
14.1971 | -105.3069
14.2007 | -105.311
14.2044 | -105.3192
14.2081 | -105.3302
14.2118 | -105.3389
14.2155 | -105.3463
14.2192 | -105.3561
14.2229 | -105.3676
14.2266 | -105.3756
14.2303 | -105.3855
14.2339 | -105.3954
14.2376 | -105.4062
14.2413 | -105.4139
14.245 | -105.4198
14.2487 | -105.4209
14.2524 | -105.4181
14.2561 | -105.4177
14.2598 | -105.4171
14.2635 | -105.4123
14.2671 | -105.4066
14.2708 | -105.4022
14.2745 | -105.3974
14.2782 | -105.3865
14.2819 | -105.3769
14.2856 | -105.3651
14.2893 | -105.3526
14.293 | -105.3424
14.2966 | -105.3351
14.3003 | -105.3311
14.304 | -105.3236
14.3077 | -105.3137
14.3114 | -105.3058
14.3151 | -105.3017
14.3188 | -105.2997
14.3225 | -105.2976
14.3262 | -105.2992
14.3298 | -105.3029
14.3335 | -105.3056
14.3372 | -105.3093
14.3409 | -105.3127
14.3446 | -105.3161
14.3483 | -105.3172
14.352 | -105.3217
14.3557 | -105.3288
14.3594 | -105.3372
14.363 | -105.3436
14.3667 | -105.3456
14.3704 | -105.3482
14.3741 | -105.3547
14.3778 | -105.3585
14.3815 | -105.3606
14.3852 | -105.3606
14.3889 | -105.3607
14.3925 | -105.3616
14.3962 | -105.3657
14.3999 | -105.3693
14.4036 | -105.3705
14.4073 | -105.3747
14.411 | -105.3768
14.4147 | -105.3788
14.4184 | -105.3821
14.4221 | -105.3871
14.4257 | -105.3927
14.4294 | -105.3966
14.4331 | -105.3992
14.4368 | -105.4013
14.4405 | -105.4037
14.4442 | -105.4048
14.4479 | -105.4063
14.4516 | -105.408
14.4553 | -105.4101
14.4589 | -105.4141
14.4626 | -105.4137
14.4663 | -105.4105
14.47 | -105.4125
14.4737 | -105.4162
14.4774 | -105.417
14.4811 | -105.4138
14.4848 | -105.4127
14.4885 | -105.4111
14.4921 | -105.4072
14.4958 | -105.4012
14.4995 | -105.3948
14.5032 | -105.3872
14.5069 | -105.3781
14.5106 | -105.3724
14.5143 | -105.3663
14.518 | -105.3593
14.5216 | -105.3502
14.5253 | -105.3411
14.529 | -105.3302
14.5327 | -105.3199
14.5364 | -105.3103
14.5401 | -105.2991
14.5438 | -105.2882
14.5475 | -105.2772
14.5512 | -105.2651
14.5548 | -105.2567
14.5585 | -105.2447
14.5622 | -105.231
14.5659 | -105.217
14.5696 | -105.206
14.5733 | -105.1953
14.577 | -105.1827
14.5807 | -105.1729
14.5844 | -105.1573
14.588 | -105.1432
14.5917 | -105.1342
14.5954 | -105.1293
14.5991 | -105.1239
14.6028 | -105.1175
14.6065 | -105.11
14.6102 | -105.1055
14.6139 | -105.1035
14.6175 | -105.0986
14.6212 | -105.096
14.6249 | -105.0912
14.6286 | -105.0886
14.6323 | -105.0892
14.636 | -105.0885
14.6397 | -105.0892
14.6434 | -105.0933
14.6471 | -105.0955
14.6507 | -105.0984
14.6544 | -105.1033
14.6581 | -105.1065
14.6618 | -105.1068
14.6655 | -105.1087
14.6692 | -105.1132
14.6729 | -105.1181
14.6766 | -105.1255
14.6803 | -105.1321
14.6839 | -105.1365
14.6876 | -105.1426
14.6913 | -105.1516
14.695 | -105.1619
14.6987 | -105.1726
14.7024 | -105.1834
14.7061 | -105.1946
14.7098 | -105.2017
14.7134 | -105.2103
14.7171 | -105.2215
14.7208 | -105.2314
14.7245 | -105.2387
14.7282 | -105.2461
14.7319 | -105.2532
14.7356 | -105.2585
14.7393 | -105.2641
14.743 | -105.267
14.7466 | -105.2675
14.7503 | -105.2705
14.754 | -105.2712
14.7577 | -105.2703
14.7614 | -105.2692
14.7651 | -105.2659
14.7688 | -105.2647
14.7725 | -105.264
14.7762 | -105.2646
14.7798 | -105.2632
14.7835 | -105.2604
14.7872 | -105.2615
14.7909 | -105.2631
14.7946 | -105.2659
14.7983 | -105.2692
14.802 | -105.2723
14.8057 | -105.2749
14.8094 | -105.2772
14.813 | -105.2802
14.8167 | -105.2811
14.8204 | -105.2804
14.8241 | -105.282
14.8278 | -105.287
14.8315 | -105.2908
14.8352 | -105.296
14.8389 | -105.3001
14.8425 | -105.2999
14.8462 | -105.2997
14.8499 | -105.3002
14.8536 | -105.2998
14.8573 | -105.2984
14.861 | -105.2941
14.8647 | -105.2885
14.8684 | -105.2846
14.8721 | -105.2822
14.8757 | -105.2751
14.8794 | -105.2641
14.8831 | -105.2563
14.8868 | -105.2475
14.8905 | -105.2381
14.8942 | -105.2301
14.8979 | -105.2217
14.9016 | -105.2152
14.9053 | -105.2096
14.9089 | -105.2035
14.9126 | -105.1955
14.9163 | -105.19
14.92 | -105.1858
14.9237 | -105.1848
14.9274 | -105.182
14.9311 | -105.1779
14.9348 | -105.1752
14.9384 | -105.1774
14.9421 | -105.1783
14.9458 | -105.1794
14.9495 | -105.1847
14.9532 | -105.1887
14.9569 | -105.1936
14.9606 | -105.1977
14.9643 | -105.2021
14.968 | -105.209
14.9716 | -105.2179
14.9753 | -105.226
14.979 | -105.2326
14.9827 | -105.2339
14.9864 | -105.2354
14.9901 | -105.2409
14.9938 | -105.2454
14.9975 | -105.2497
15.0012 | -105.2528
15.0048 | -105.2521
15.0085 | -105.2499
15.0122 | -105.2475
15.0159 | -105.2436
15.0196 | -105.238
15.0233 | -105.2338
15.027 | -105.2298
15.0307 | -105.2255
15.0343 | -105.2204
15.038 | -105.2129
15.0417 | -105.2043
15.0454 | -105.1949
15.0491 | -105.1872
15.0528 | -105.1795
15.0565 | -105.1741
15.0602 | -105.1681
15.0639 | -105.1624
15.0675 | -105.1589
15.0712 | -105.1597
15.0749 | -105.1626
15.0786 | -105.1667
15.0823 | -105.1713
15.086 | -105.1749
15.0897 | -105.1823
15.0934 | -105.1938
15.0971 | -105.2043
15.1007 | -105.2182
15.1044 | -105.2324
15.1081 | -105.2463
15.1118 | -105.2605
15.1155 | -105.2729
15.1192 | -105.2906
15.1229 | -105.3084
15.1266 | -105.3239
15.1303 | -105.3368
15.1339 | -105.3472
15.1376 | -105.3582
15.1413 | -105.3679
15.145 | -105.3793
15.1487 | -105.3868
15.1524 | -105.3891
15.1561 | -105.3892
15.1598 | -105.386
15.1634 | -105.3771
15.1671 | -105.3677
15.1708 | -105.3605
15.1745 | -105.3507
15.1782 | -105.3391
15.1819 | -105.3252
15.1856 | -105.3088
15.1893 | -105.289
15.193 | -105.2713
15.1966 | -105.2581
15.2003 | -105.2477
15.204 | -105.2343
15.2077 | -105.2174
15.2114 | -105.2027
15.2151 | -105.1897
15.2188 | -105.1773
15.2225 | -105.1682
15.2262 | -105.1632
15.2298 | -105.1595
15.2335 | -105.1549
15.2372 | -105.1542
15.2409 | -105.1552
15.2446 | -105.1566
15.2483 | -105.1616
15.252 | -105.1685
15.2557 | -105.1786
15.2593 | -105.1894
15.263 | -105.1978
15.2667 | -105.2065
15.2704 | -105.2189
15.2741 | -105.2303
15.2778 | -105.2414
15.2815 | -105.2528
15.2852 | -105.2645
15.2889 | -105.2744
15.2925 | -105.2854
15.2962 | -105.2947
15.2999 | -105.3039
15.3036 | -105.3114
15.3073 | -105.3154
15.311 | -105.319
15.3147 | -105.3215
15.3184 | -105.3229
15.3221 | -105.3225
15.3257 | -105.3226
15.3294 | -105.323
15.3331 | -105.3219
15.3368 | -105.3163
15.3405 | -105.3099
15.3442 | -105.3046
15.3479 | -105.3026
15.3516 | -105.3017
15.3552 | -105.297
15.3589 | -105.2888
15.3626 | -105.2799
15.3663 | -105.2734
15.37 | -105.2685
15.3737 | -105.2625
15.3774 | -105.2601
15.3811 | -105.2598
15.3848 | -105.2566
15.3884 | -105.2504
15.3921 | -105.2454
15.3958 | -105.2428
15.3995 | -105.2414
15.4032 | -105.2411
15.4069 | -105.2432
15.4106 | -105.244
15.4143 | -105.2423
15.418 | -105.241
15.4216 | -105.2427
15.4253 | -105.2467
15.429 | -105.2501
15.4327 | -105.249
15.4364 | -105.2505
15.4401 | -105.2521
15.4438 | -105.2545
15.4475 | -105.2569
15.4512 | -105.2601
15.4548 | -105.2632
15.4585 | -105.2641
15.4622 | -105.2656
15.4659 | -105.2683
15.4696 | -105.2702
15.4733 | -105.272
15.477 | -105.2734
15.4807 | -105.277
15.4843 | -105.285
15.488 | -105.2913
15.4917 | -105.2952
15.4954 | -105.3027
15.4991 | -105.3124
15.5028 | -105.3211
15.5065 | -105.326
15.5102 | -105.3313
15.5139 | -105.3369
15.5175 | -105.3453
15.5212 | -105.3548
15.5249 | -105.3661
15.5286 | -105.3767
15.5323 | -105.3846
15.536 | -105.3904
15.5397 | -105.3937
15.5434 | -105.3992
15.5471 | -105.404
15.5507 | -105.4073
15.5544 | -105.4077
15.5581 | -105.4072
15.5618 | -105.4091
15.5655 | -105.4102
15.5692 | -105.4056
15.5729 | -105.3971
15.5766 | -105.3946
15.5802 | -105.3937
15.5839 | -105.3901
15.5876 | -105.3861
15.5913 | -105.3795
15.595 | -105.3708
15.5987 | -105.3623
15.6024 | -105.3514
15.6061 | -105.3409
15.6098 | -105.3332
15.6134 | -105.3254
15.6171 | -105.3159
15.6208 | -105.3053
15.6245 | -105.2949
15.6282 | -105.285
15.6319 | -105.2766
15.6356 | -105.2688
15.6393 | -105.2618
15.643 | -105.2543
15.6466 | -105.2475
15.6503 | -105.2421
15.654 | -105.2402
15.6577 | -105.2417
15.6614 | -105.2432
15.6651 | -105.2462
15.6688 | -105.2505
15.6725 | -105.2557
15.6761 | -105.2618
15.6798 | -105.2712
15.6835 | -105.2822
15.6872 | -105.2933
15.6909 | -105.3039
15.6946 | -105.3139
15.6983 | -105.327
15.702 | -105.343
15.7057 | -105.358
15.7093 | -105.3721
15.713 | -105.3856
15.7167 | -105.4018
15.7204 | -105.4148
15.7241 | -105.4296
15.7278 | -105.4439
15.7315 | -105.4586
15.7352 | -105.4712
15.7389 | -105.4812
15.7425 | -105.4904
15.7462 | -105.4971
15.7499 | -105.5046
15.7536 | -105.5105
15.7573 | -105.5138
15.761 | -105.5176
15.7647 | -105.5223
15.7684 | -105.5233
15.7721 | -105.5208
15.7757 | -105.5165
15.7794 | -105.5145
15.7831 | -105.5098
15.7868 | -105.4995
15.7905 | -105.4887
15.7942 | -105.4828
15.7979 | -105.4769
15.8016 | -105.4622
15.8052 | -105.4475
15.8089 | -105.4328
15.8126 | -105.4186
15.8163 | -105.4057
15.82 | -105.3918
15.8237 | -105.3785
15.8274 | -105.3666
15.8311 | -105.3526
15.8348 | -105.3354
15.8384 | -105.3238
15.8421 | -105.3135
15.8458 | -105.3027
15.8495 | -105.2887
15.8532 | -105.2763
15.8569 | -105.2668
15.8606 | -105.258
15.8643 | -105.2517
15.868 | -105.2466
15.8716 | -105.2408
15.8753 | -105.2371
15.879 | -105.2338
15.8827 | -105.229
15.8864 | -105.2236
15.8901 | -105.2194
15.8938 | -105.213
15.8975 | -105.2084
15.9011 | -105.2066
15.9048 | -105.2047
15.9085 | -105.2032
15.9122 | -105.1995
15.9159 | -105.1966
15.9196 | -105.1938
15.9233 | -105.19
15.927 | -105.1857
15.9307 | -105.1804
15.9343 | -105.1723
15.938 | -105.1674
15.9417 | -105.1645
15.9454 | -105.162
15.9491 | -105.1579
15.9528 | -105.1552
15.9565 | -105.1522
15.9602 | -105.1498
15.9639 | -105.1482
15.9675 | -105.1469
15.9712 | -105.1448
15.9749 | -105.142
15.9786 | -105.1368
15.9823 | -105.1328
15.986 | -105.1295
15.9897 | -105.1254
15.9934 | -105.1178
15.997 | -105.1093
16.0007 | -105.1038
16.0044 | -105.0972
16.0081 | -105.088
16.0118 | -105.0801
16.0155 | -105.074
16.0192 | -105.0664
16.0229 | -105.057
16.0266 | -105.0492
16.0302 | -105.042
16.0339 | -105.0363
16.0376 | -105.0323
16.0413 | -105.0281
16.045 | -105.0246
16.0487 | -105.0223
16.0524 | -105.019
16.0561 | -105.0177
16.0598 | -105.0185
16.0634 | -105.0221
16.0671 | -105.0222
16.0708 | -105.0242
16.0745 | -105.0277
16.0782 | -105.0343
16.0819 | -105.0397
16.0856 | -105.0435
16.0893 | -105.0472
16.0929 | -105.0492
16.0966 | -105.0534
16.1003 | -105.0587
16.104 | -105.0634
16.1077 | -105.0696
16.1114 | -105.0779
16.1151 | -105.0862
16.1188 | -105.0927
16.1225 | -105.101
16.1261 | -105.106
16.1298 | -105.1082
16.1335 | -105.1126
16.1372 | -105.1176
16.1409 | -105.1246
16.1446 | -105.1299
16.1483 | -105.1344
16.152 | -105.1376
16.1557 | -105.1405
16.1593 | -105.1461
16.163 | -105.1541
16.1667 | -105.159
16.1704 | -105.1588
16.1741 | -105.1602
16.1778 | -105.162
16.1815 | -105.1651
16.1852 | -105.1706
16.1889 | -105.176
16.1925 | -105.1788
16.1962 | -105.1836
16.1999 | -105.1868
16.2036 | -105.1908
16.2073 | -105.1956
16.211 | -105.2015
16.2147 | -105.2077
16.2184 | -105.2141
16.222 | -105.2187
16.2257 | -105.2256
16.2294 | -105.2309
16.2331 | -105.2317
16.2368 | -105.2326
16.2405 | -105.2313
16.2442 | -105.229
16.2479 | -105.2248
16.2516 | -105.2192
16.2552 | -105.2103
16.2589 | -105.2008
16.2626 | -105.1942
16.2663 | -105.1847
16.27 | -105.1732
16.2737 | -105.1614
16.2774 | -105.15
16.2811 | -105.1355
16.2848 | -105.1204
16.2884 | -105.1074
16.2921 | -105.0932
16.2958 | -105.0795
16.2995 | -105.0678
16.3032 | -105.0566
16.3069 | -105.0459
16.3106 | -105.0392
16.3143 | -105.0334
16.3179 | -105.0273
16.3216 | -105.0218
16.3253 | -105.0181
16.329 | -105.0174
16.3327 | -105.018
16.3364 | -105.0188
16.3401 | -105.0209
16.3438 | -105.0252
16.3475 | -105.0321
16.3511 | -105.0407
16.3548 | -105.0501
16.3585 | -105.062
16.3622 | -105.076
16.3659 | -105.0893
16.3696 | -105.1025
16.3733 | -105.1151
16.377 | -105.1292
16.3807 | -105.144
16.3843 | -105.1597
16.388 | -105.1735
16.3917 | -105.1848
16.3954 | -105.1976
16.3991 | -105.2087
16.4028 | -105.2195
16.4065 | -105.2313
16.4102 | -105.2425
16.4139 | -105.2524
16.4175 | -105.2601
16.4212 | -105.2682
16.4249 | -105.2764
16.4286 | -105.2833
16.4323 | -105.2922
16.436 | -105.2987
16.4397 | -105.3037
16.4434 | -105.3056
16.447 | -105.3061
16.4507 | -105.3062
16.4544 | -105.3093
16.4581 | -105.3139
16.4618 | -105.3208
16.4655 | -105.3275
16.4692 | -105.3328
16.4729 | -105.3395
16.4766 | -105.3465
16.4802 | -105.3542
16.4839 | -105.3622
16.4876 | -105.3689
16.4913 | -105.3721
16.495 | -105.3793
16.4987 | -105.3841
16.5024 | -105.3896
16.5061 | -105.3966
16.5098 | -105.4048
16.5134 | -105.4128
16.5171 | -105.4235
16.5208 | -105.4301
16.5245 | -105.437
16.5282 | -105.4456
16.5319 | -105.4522
16.5356 | -105.4592
16.5393 | -105.4645
16.5429 | -105.4685
16.5466 | -105.4718
16.5503 | -105.4773
16.554 | -105.4845
16.5577 | -105.4939
16.5614 | -105.5026
16.5651 | -105.5077
16.5688 | -105.5135
16.5725 | -105.5219
16.5761 | -105.5298
16.5798 | -105.5372
16.5835 | -105.5442
16.5872 | -105.5532
16.5909 | -105.5643
16.5946 | -105.5742
16.5983 | -105.5816
16.602 | -105.5853
16.6057 | -105.5914
16.6093 | -105.6007
16.613 | -105.6093
16.6167 | -105.6177
16.6204 | -105.6253
16.6241 | -105.634
16.6278 | -105.6459
16.6315 | -105.6556
16.6352 | -105.6659
16.6388 | -105.6767
16.6425 | -105.6842
16.6462 | -105.6908
16.6499 | -105.6982
16.6536 | -105.7037
16.6573 | -105.7064
16.661 | -105.7092
16.6647 | -105.7123
16.6684 | -105.7156
16.672 | -105.7158
16.6757 | -105.7132
16.6794 | -105.7134
16.6831 | -105.7121
16.6868 | -105.71
16.6905 | -105.7046
16.6942 | -105.6972
16.6979 | -105.6884
16.7016 | -105.6814
16.7052 | -105.6749
16.7089 | -105.6667
16.7126 | -105.6576
16.7163 | -105.6492
16.72 | -105.6417
16.7237 | -105.6332
16.7274 | -105.6242
16.7311 | -105.612
16.7347 | -105.5988
16.7384 | -105.5858
16.7421 | -105.573
16.7458 | -105.5617
16.7495 | -105.5502
16.7532 | -105.539
16.7569 | -105.5297
16.7606 | -105.5237
16.7643 | -105.5142
16.7679 | -105.5064
16.7716 | -105.5016
16.7753 | -105.4975
16.779 | -105.4953
16.7827 | -105.4967
16.7864 | -105.4971
16.7901 | -105.4963
16.7938 | -105.4947
16.7975 | -105.4942
16.8011 | -105.4962
16.8048 | -105.4983
16.8085 | -105.5023
16.8122 | -105.5068
16.8159 | -105.5114
16.8196 | -105.5161
16.8233 | -105.5203
16.827 | -105.5243
16.8307 | -105.529
16.8343 | -105.5362
16.838 | -105.5462
16.8417 | -105.5591
16.8454 | -105.5669
16.8491 | -105.572
16.8528 | -105.5749
16.8565 | -105.5777
16.8602 | -105.581
16.8638 | -105.5849
16.8675 | -105.5875
16.8712 | -105.59
16.8749 | -105.5927
16.8786 | -105.5967
16.8823 | -105.5999
16.886 | -105.6014
16.8897 | -105.6029
16.8934 | -105.6067
16.897 | -105.6104
16.9007 | -105.6111
16.9044 | -105.6105
16.9081 | -105.6096
16.9118 | -105.6098
16.9155 | -105.61
16.9192 | -105.6097
16.9229 | -105.6084
16.9266 | -105.6054
16.9302 | -105.6028
16.9339 | -105.5989
16.9376 | -105.5966
16.9413 | -105.5948
16.945 | -105.5915
16.9487 | -105.5892
16.9524 | -105.5858
16.9561 | -105.5827
16.9597 | -105.5808
16.9634 | -105.5795
16.9671 | -105.579
16.9708 | -105.5795
16.9745 | -105.5774
16.9782 | -105.575
16.9819 | -105.5766
16.9856 | -105.5772
16.9893 | -105.5758
16.9929 | -105.5745
16.9966 | -105.574
17.0003 | -105.5745
17.004 | -105.5742
17.0077 | -105.5724
17.0114 | -105.5707
17.0151 | -105.5698
17.0188 | -105.5683
17.0225 | -105.5643
17.0261 | -105.5629
17.0298 | -105.5607
17.0335 | -105.5572
17.0372 | -105.5546
17.0409 | -105.5531
17.0446 | -105.5522
17.0483 | -105.5502
17.052 | -105.5469
17.0556 | -105.5445
17.0593 | -105.5419
17.063 | -105.5395
17.0667 | -105.5373
17.0704 | -105.5357
17.0741 | -105.5341
17.0778 | -105.5308
17.0815 | -105.5255
17.0852 | -105.5217
17.0888 | -105.5214
17.0925 | -105.5217
17.0962 | -105.5255
17.0999 | -105.528
17.1036 | -105.5284
17.1073 | -105.5308
17.111 | -105.5325
17.1147 | -105.5346
17.1184 | -105.5371
17.122 | -105.5393
17.1257 | -105.5433
17.1294 | -105.5495
17.1331 | -105.5566
17.1368 | -105.5633
17.1405 | -105.5726
17.1442 | -105.5833
17.1479 | -105.5932
17.1516 | -105.6019
17.1552 | -105.611
17.1589 | -105.6204
17.1626 | -105.6295
17.1663 | -105.6372
17.17 | -105.6455
17.1737 | -105.6538
17.1774 | -105.6636
17.1811 | -105.6719
17.1847 | -105.6777
17.1884 | -105.6853
17.1921 | -105.6941
17.1958 | -105.703
17.1995 | -105.7096
17.2032 | -105.714
17.2069 | -105.7216
17.2106 | -105.7286
17.2143 | -105.7311
17.2179 | -105.7337
17.2216 | -105.737
17.2253 | -105.7355
17.229 | -105.7326
17.2327 | -105.7305
17.2364 | -105.7276
17.2401 | -105.7226
17.2438 | -105.7154
17.2475 | -105.7092
17.2511 | -105.7033
17.2548 | -105.6943
17.2585 | -105.6828
17.2622 | -105.6736
17.2659 | -105.6658
17.2696 | -105.6554
17.2733 | -105.6422
17.277 | -105.629
17.2806 | -105.6191
17.2843 | -105.6083
17.288 | -105.5963
17.2917 | -105.5852
17.2954 | -105.5755
17.2991 | -105.5658
17.3028 | -105.5544
17.3065 | -105.5447
17.3102 | -105.5362
17.3138 | -105.5291
17.3175 | -105.5212
17.3212 | -105.5129
17.3249 | -105.5077
17.3286 | -105.5028
17.3323 | -105.498
17.336 | -105.4952
17.3397 | -105.4923
17.3434 | -105.4884
17.347 | -105.4877
17.3507 | -105.486
17.3544 | -105.484
17.3581 | -105.4855
17.3618 | -105.4874
17.3655 | -105.4897
17.3692 | -105.4918
17.3729 | -105.4939
17.3766 | -105.4955
17.3802 | -105.4963
17.3839 | -105.4995
17.3876 | -105.5023
17.3913 | -105.5039
17.395 | -105.5059
17.3987 | -105.5078
17.4024 | -105.5104
17.4061 | -105.5135
17.4097 | -105.5146
17.4134 | -105.514
17.4171 | -105.5123
17.4208 | -105.5119
17.4245 | -105.5082
17.4282 | -105.5025
17.4319 | -105.4993
17.4356 | -105.4949
17.4393 | -105.4866
17.4429 | -105.4776
17.4466 | -105.469
17.4503 | -105.4602
17.454 | -105.4469
17.4577 | -105.4322
17.4614 | -105.4193
17.4651 | -105.4065
17.4688 | -105.3912
17.4725 | -105.3765
17.4761 | -105.3617
17.4798 | -105.346
17.4835 | -105.3309
17.4872 | -105.3174
17.4909 | -105.3034
17.4946 | -105.2899
17.4983 | -105.2764
17.502 | -105.2633
17.5056 | -105.252
17.5093 | -105.2394
17.513 | -105.2271
17.5167 | -105.2173
17.5204 | -105.2092
17.5241 | -105.2
17.5278 | -105.1909
17.5315 | -105.1821
17.5352 | -105.1756
17.5388 | -105.1688
17.5425 | -105.1646
17.5462 | -105.1627
17.5499 | -105.1621
17.5536 | -105.162
17.5573 | -105.1692
17.561 | -105.17
17.5647 | -105.1712
17.5684 | -105.174
17.572 | -105.1772
17.5757 | -105.1828
17.5794 | -105.1909
17.5831 | -105.1983
17.5868 | -105.2032
17.5905 | -105.2102
17.5942 | -105.2193
17.5979 | -105.2275
17.6015 | -105.238
17.6052 | -105.2487
17.6089 | -105.2569
17.6126 | -105.2664
17.6163 | -105.2771
17.62 | -105.288
17.6237 | -105.2981
17.6274 | -105.3059
17.6311 | -105.314
17.6347 | -105.3213
17.6384 | -105.3285
17.6421 | -105.3362
17.6458 | -105.3419
17.6495 | -105.3472
17.6532 | -105.3508
17.6569 | -105.3529
17.6606 | -105.3551
17.6643 | -105.3587
17.6679 | -105.3617
17.6716 | -105.3616
17.6753 | -105.36
17.679 | -105.3594
17.6827 | -105.3564
17.6864 | -105.3526
17.6901 | -105.3507
17.6938 | -105.3491
17.6974 | -105.3487
17.7011 | -105.3471
17.7048 | -105.345
17.7085 | -105.3439
17.7122 | -105.3422
17.7159 | -105.3402
17.7196 | -105.3395
17.7233 | -105.3407
17.727 | -105.3412
17.7306 | -105.3416
17.7343 | -105.3406
17.738 | -105.3395
17.7417 | -105.3405
17.7454 | -105.3432
17.7491 | -105.3475
17.7528 | -105.3531
17.7565 | -105.3571
17.7602 | -105.3623
17.7638 | -105.3683
17.7675 | -105.3708
17.7712 | -105.3778
17.7749 | -105.3878
17.7786 | -105.3978
17.7823 | -105.4086
17.786 | -105.4161
17.7897 | -105.4244
17.7934 | -105.4333
17.797 | -105.4395
17.8007 | -105.4463
17.8044 | -105.4537
17.8081 | -105.4621
17.8118 | -105.4718
17.8155 | -105.481
17.8192 | -105.4885
17.8229 | -105.4941
17.8265 | -105.4999
17.8302 | -105.5025
17.8339 | -105.5056
17.8376 | -105.5104
17.8413 | -105.5155
17.845 | -105.5186
17.8487 | -105.5198
17.8524 | -105.5243
17.8561 | -105.5287
17.8597 | -105.5301
17.8634 | -105.5311
17.8671 | -105.5316
17.8708 | -105.5294
17.8745 | -105.5268
17.8782 | -105.5226
17.8819 | -105.5195
17.8856 | -105.5141
17.8893 | -105.5096
17.8929 | -105.5041
17.8966 | -105.4992
17.9003 | -105.4952
17.904 | -105.4909
17.9077 | -105.4867
17.9114 | -105.4824
17.9151 | -105.4759
17.9188 | -105.4691
17.9224 | -105.4629
17.9261 | -105.4585
17.9298 | -105.4559
17.9335 | -105.4507
17.9372 | -105.4427
17.9409 | -105.4343
17.9446 | -105.4269
17.9483 | -105.421
17.952 | -105.415
17.9556 | -105.4108
17.9593 | -105.4063
17.963 | -105.4015
17.9667 | -105.3968
17.9704 | -105.3933
17.9741 | -105.3912
17.9778 | -105.3872
17.9815 | -105.3842
17.9852 | -105.3829
17.9888 | -105.3839
17.9925 | -105.3836
17.9962 | -105.3814
17.9999 | -105.3779
18.0036 | -105.3733
18.0073 | -105.371
18.011 | -105.3695
18.0147 | -105.3697
18.0184 | -105.3705
18.022 | -105.3718
18.0257 | -105.3714
18.0294 | -105.3699
18.0331 | -105.3703
18.0368 | -105.3726
18.0405 | -105.3747
18.0442 | -105.3752
18.0479 | -105.3756
18.0515 | -105.3766
18.0552 | -105.3796
18.0589 | -105.3828
18.0626 | -105.3834
18.0663 | -105.3828
18.07 | -105.3848
18.0737 | -105.3873
18.0774 | -105.3903
18.0811 | -105.3922
18.0847 | -105.3955
18.0884 | -105.3993
18.0921 | -105.4024
18.0958 | -105.4026
18.0995 | -105.4027
18.1032 | -105.4056
18.1069 | -105.4095
18.1106 | -105.4103
18.1143 | -105.4114
18.1179 | -105.4156
18.1216 | -105.4212
18.1253 | -105.4238
18.129 | -105.4274
18.1327 | -105.4315
18.1364 | -105.4382
18.1401 | -105.4443
18.1438 | -105.4484
18.1474 | -105.4488
18.1511 | -105.4536
18.1548 | -105.4573
18.1585 | -105.4598
18.1622 | -105.4608
18.1659 | -105.4623
18.1696 | -105.4634
18.1733 | -105.465
18.177 | -105.4638
18.1806 | -105.4615
18.1843 | -105.4599
18.188 | -105.4571
18.1917 | -105.4558
18.1954 | -105.4579
18.1991 | -105.4588
18.2028 | -105.4572
18.2065 | -105.4539
18.2102 | -105.4482
18.2138 | -105.4416
18.2175 | -105.436
18.2212 | -105.4288
18.2249 | -105.422
18.2286 | -105.414
18.2323 | -105.404
18.236 | -105.3934
18.2397 | -105.3856
18.2433 | -105.3775
18.247 | -105.3679
18.2507 | -105.3584
18.2544 | -105.3505
18.2581 | -105.3419
18.2618 | -105.3352
18.2655 | -105.327
18.2692 | -105.3179
18.2729 | -105.3091
18.2765 | -105.3008
18.2802 | -105.2925
18.2839 | -105.2856
18.2876 | -105.2789
18.2913 | -105.2713
18.295 | -105.2648
18.2987 | -105.2609
18.3024 | -105.258
18.3061 | -105.2553
18.3097 | -105.2543
18.3134 | -105.2525
18.3171 | -105.2498
18.3208 | -105.2487
18.3245 | -105.2514
18.3282 | -105.2533
18.3319 | -105.2535
18.3356 | -105.2559
18.3392 | -105.2577
18.3429 | -105.2607
18.3466 | -105.2637
18.3503 | -105.2666
18.354 | -105.2708
18.3577 | -105.274
18.3614 | -105.2797
18.3651 | -105.2832
18.3688 | -105.2833
18.3724 | -105.2842
18.3761 | -105.2854
18.3798 | -105.2853
18.3835 | -105.2852
18.3872 | -105.2873
18.3909 | -105.2875
18.3946 | -105.2857
18.3983 | -105.2839
18.402 | -105.2812
18.4056 | -105.2776
18.4093 | -105.2754
18.413 | -105.2715
18.4167 | -105.2679
18.4204 | -105.2611
18.4241 | -105.2546
18.4278 | -105.2506
18.4315 | -105.2472
18.4352 | -105.2439
18.4388 | -105.242
18.4425 | -105.2395
18.4462 | -105.2349
18.4499 | -105.2309
18.4536 | -105.2271
18.4573 | -105.2204
18.461 | -105.2136
18.4647 | -105.207
18.4683 | -105.2014
18.472 | -105.1989
18.4757 | -105.1968
18.4794 | -105.195
18.4831 | -105.1936
18.4868 | -105.1891
18.4905 | -105.1845
18.4942 | -105.1808
18.4979 | -105.179
18.5015 | -105.1751
18.5052 | -105.1698
18.5089 | -105.1647
18.5126 | -105.1596
18.5163 | -105.1528
18.52 | -105.1465
18.5237 | -105.1379
18.5274 | -105.1299
18.5311 | -105.124
18.5347 | -105.1144
18.5384 | -105.1047
18.5421 | -105.0988
18.5458 | -105.0894
18.5495 | -105.0792
18.5532 | -105.0696
18.5569 | -105.0606
18.5606 | -105.0513
18.5642 | -105.042
18.5679 | -105.0324
18.5716 | -105.0239
18.5753 | -105.0137
18.579 | -105.0036
18.5827 | -104.9952
18.5864 | -104.9844
18.5901 | -104.9739
18.5938 | -104.9671
18.5974 | -104.9608
18.6011 | -104.955
18.6048 | -104.9466
18.6085 | -104.9362
18.6122 | -104.9269
18.6159 | -104.9195
18.6196 | -104.9137
18.6233 | -104.9074
18.627 | -104.8995
18.6306 | -104.8929
18.6343 | -104.8887
18.638 | -104.8842
18.6417 | -104.8823
18.6454 | -104.8781
18.6491 | -104.8729
18.6528 | -104.8692
18.6565 | -104.8668
18.6602 | -104.8656
18.6638 | -104.863
18.6675 | -104.8632
18.6712 | -104.8645
18.6749 | -104.8621
18.6786 | -104.8564
18.6823 | -104.8527
18.686 | -104.8536
18.6897 | -104.8542
18.6933 | -104.8534
18.697 | -104.8533
18.7007 | -104.8529
18.7044 | -104.8528
18.7081 | -104.8521
18.7118 | -104.8517
18.7155 | -104.8517
18.7192 | -104.8504
18.7229 | -104.8498
18.7265 | -104.8515
18.7302 | -104.8535
18.7339 | -104.8544
18.7376 | -104.8543
18.7413 | -104.8548
18.745 | -104.8558
18.7487 | -104.8566
18.7524 | -104.8587
18.7561 | -104.8581
18.7597 | -104.8548
18.7634 | -104.8528
18.7671 | -104.8524
18.7708 | -104.8523
18.7745 | -104.8505
18.7782 | -104.8499
18.7819 | -104.8493
18.7856 | -104.8467
18.7892 | -104.8441
18.7929 | -104.841
18.7966 | -104.8387
18.8003 | -104.836
18.804 | -104.8314
18.8077 | -104.8267
18.8114 | -104.8224
18.8151 | -104.8178
18.8188 | -104.8135
18.8224 | -104.8082
18.8261 | -104.8021
18.8298 | -104.7949
18.8335 | -104.7902
18.8372 | -104.7862
18.8409 | -104.7808
18.8446 | -104.7762
18.8483 | -104.7708
18.852 | -104.7652
18.8556 | -104.7608
18.8593 | -104.7556
18.863 | -104.7496
18.8667 | -104.7465
18.8704 | -104.7433
18.8741 | -104.7399
18.8778 | -104.7364
18.8815 | -104.7317
18.8851 | -104.7256
18.8888 | -104.7208
18.8925 | -104.7187
18.8962 | -104.7139
18.8999 | -104.7089
18.9036 | -104.7021
18.9073 | -104.6942
18.911 | -104.685
18.9147 | -104.6785
18.9183 | -104.671
18.922 | -104.6627
18.9257 | -104.6496
18.9294 | -104.6384
18.9331 | -104.6291
18.9368 | -104.618
18.9405 | -104.6062
18.9442 | -104.5967
18.9479 | -104.5874
18.9515 | -104.5744
18.9552 | -104.5599
18.9589 | -104.5472
18.9626 | -104.53
18.9663 | -104.5126
18.97 | -104.496
18.9737 | -104.4789
18.9774 | -104.4622
18.981 | -104.4454
18.9847 | -104.4289
18.9884 | -104.4115
18.9921 | -104.3929
18.9958 | -104.3736
18.9995 | -104.3537
19.0032 | -104.3325
19.0069 | -104.3132
19.0106 | -104.2946
19.0142 | -104.2749
19.0179 | -104.2546
19.0216 | -104.2358
19.0253 | -104.2209
19.029 | -104.2085
19.0327 | -104.1936
19.0364 | -104.1789
19.0401 | -104.1646
19.0438 | -104.1527
19.0474 | -104.1415
19.0511 | -104.1313
19.0548 | -104.1216
19.0585 | -104.1112
19.0622 | -104.102
19.0659 | -104.0917
19.0696 | -104.0849
19.0733 | -104.0804
19.077 | -104.0758
19.0806 | -104.0712
19.0843 | -104.0675
19.088 | -104.0635
19.0917 | -104.0611
19.0954 | -104.0599
19.0991 | -104.0598
19.1028 | -104.0582
19.1065 | -104.0575
19.1101 | -104.0578
19.1138 | -104.058
19.1175 | -104.0588
19.1212 | -104.0607
19.1249 | -104.0602
19.1286 | -104.0596
19.1323 | -104.0608
19.136 | -104.063
19.1397 | -104.0645
19.1433 | -104.0663
19.147 | -104.0683
19.1507 | -104.0705
19.1544 | -104.0676
19.1581 | -104.0681
19.1618 | -104.068
19.1655 | -104.0652
19.1692 | -104.0649
19.1729 | -104.0636
19.1765 | -104.0637
19.1802 | -104.0613
19.1839 | -104.0555
19.1876 | -104.0498
19.1913 | -104.0459
19.195 | -104.0401
19.1987 | -104.0369
19.2024 | -104.029
19.206 | -104.0209
19.2097 | -104.0131
19.2134 | -104.0059
19.2171 | -103.9979
19.2208 | -103.9901
19.2245 | -103.9811
19.2282 | -103.9726
19.2319 | -103.9631
19.2356 | -103.9556
19.2392 | -103.9448
19.2429 | -103.9349
19.2466 | -103.9277
19.2503 | -103.9216
19.254 | -103.914
19.2577 | -103.9076
19.2614 | -103.903
19.2651 | -103.8989
19.2688 | -103.8939
19.2724 | -103.888
19.2761 | -103.8864
19.2798 | -103.8837
19.2835 | -103.8805
19.2872 | -103.8764
19.2909 | -103.8739
19.2946 | -103.8722
19.2983 | -103.8696
19.302 | -103.8676
19.3056 | -103.865
19.3093 | -103.8613
19.313 | -103.8609
19.3167 | -103.8612
19.3204 | -103.8609
19.3241 | -103.8603
19.3278 | -103.8621
19.3315 | -103.8631
19.3351 | -103.8629
19.3388 | -103.8639
19.3425 | -103.8632
19.3462 | -103.8619
19.3499 | -103.8589
19.3536 | -103.8578
19.3573 | -103.8568
19.361 | -103.855
19.3647 | -103.8523
19.3683 | -103.8484
19.372 | -103.8459
19.3757 | -103.8441
19.3794 | -103.8398
19.3831 | -103.8373
19.3868 | -103.8339
19.3905 | -103.8298
19.3942 | -103.8255
19.3979 | -103.8245
19.4015 | -103.8228
19.4052 | -103.8187
19.4089 | -103.8131
19.4126 | -103.8093
19.4163 | -103.805
19.42 | -103.8024
19.4237 | -103.8009
19.4274 | -103.7989
19.431 | -103.7981
19.4347 | -103.7977
19.4384 | -103.798
19.4421 | -103.7989
19.4458 | -103.7992
19.4495 | -103.801
19.4532 | -103.8047
19.4569 | -103.81
19.4606 | -103.8173
19.4642 | -103.8257
19.4679 | -103.8335
19.4716 | -103.8417
19.4753 | -103.8527
19.479 | -103.8646
19.4827 | -103.8761
19.4864 | -103.8888
19.4901 | -103.9017
19.4938 | -103.9149
19.4974 | -103.9266
19.5011 | -103.9404
19.5048 | -103.9534
19.5085 | -103.9655
19.5122 | -103.9756
19.5159 | -103.9869
19.5196 | -103.9963
19.5233 | -104.006
19.5269 | -104.0146
19.5306 | -104.0236
19.5343 | -104.0308
19.538 | -104.0366
19.5417 | -104.043
19.5454 | -104.0482
19.5491 | -104.0523
19.5528 | -104.0508
19.5565 | -104.0502
19.5601 | -104.05
19.5638 | -104.0497
19.5675 | -104.0461
19.5712 | -104.0405
19.5749 | -104.0322
19.5786 | -104.0245
19.5823 | -104.0194
19.586 | -104.0127
19.5897 | -104.0027
19.5933 | -103.9912
19.597 | -103.9813
19.6007 | -103.9706
19.6044 | -103.9591
19.6081 | -103.9473
19.6118 | -103.9357
19.6155 | -103.9225
19.6192 | -103.9095
19.6228 | -103.8969
19.6265 | -103.8867
19.6302 | -103.8752
19.6339 | -103.863
19.6376 | -103.852
19.6413 | -103.8398
19.645 | -103.8295
19.6487 | -103.8186
19.6524 | -103.8073
19.656 | -103.7977
19.6597 | -103.7875
19.6634 | -103.7813
19.6671 | -103.7761
19.6708 | -103.7709
19.6745 | -103.7655
19.6782 | -103.7633
19.6819 | -103.763
19.6856 | -103.7692
19.6892 | -103.7693
19.6929 | -103.7724
19.6966 | -103.7763
19.7003 | -103.7768
19.704 | -103.7783
19.7077 | -103.7824
19.7114 | -103.7884
19.7151 | -103.7953
19.7188 | -103.8019
19.7224 | -103.8105
19.7261 | -103.8186
19.7298 | -103.8278
19.7335 | -103.8354
19.7372 | -103.8441
19.7409 | -103.8527
19.7446 | -103.8585
19.7483 | -103.865
19.7519 | -103.8733
19.7556 | -103.8785
19.7593 | -103.8845
19.763 | -103.8911
19.7667 | -103.8968
19.7704 | -103.9032
19.7741 | -103.9084
19.7778 | -103.9157
19.7815 | -103.9224
19.7851 | -103.9259
19.7888 | -103.9304
19.7925 | -103.9376
19.7962 | -103.9424
19.7999 | -103.9469
19.8036 | -103.946
19.8073 | -103.9449
19.811 | -103.9448
19.8147 | -103.9456
19.8183 | -103.9462
19.822 | -103.9457
19.8257 | -103.9441
19.8294 | -103.9438
19.8331 | -103.9454
19.8368 | -103.9423
19.8405 | -103.9383
19.8442 | -103.934
19.8478 | -103.9317
19.8515 | -103.9292
19.8552 | -103.9272
19.8589 | -103.9227
19.8626 | -103.9182
19.8663 | -103.913
19.87 | -103.9077
19.8737 | -103.9024
19.8774 | -103.896
19.881 | -103.8911
19.8847 | -103.8861
19.8884 | -103.8816
19.8921 | -103.8768
19.8958 | -103.8728
19.8995 | -103.8675
19.9032 | -103.8608
19.9069 | -103.854
19.9106 | -103.8479
19.9142 | -103.8437
19.9179 | -103.8393
19.9216 | -103.8365
19.9253 | -103.8341
19.929 | -103.8304
19.9327 | -103.8285
19.9364 | -103.8258
19.9401 | -103.8256
19.9438 | -103.8239
19.9474 | -103.8225
19.9511 | -103.8205
19.9548 | -103.8176
19.9585 | -103.8172
19.9622 | -103.8187
19.9659 | -103.8228
19.9696 | -103.8256
19.9733 | -103.8293
19.9769 | -103.8319
19.9806 | -103.8378
19.9843 | -103.8468
19.988 | -103.8537
19.9917 | -103.8588
19.9954 | -103.864
19.9991 | -103.8691
20.0028 | -103.8736
20.0065 | -103.8784
20.0101 | -103.8863
20.0138 | -103.8934
20.0175 | -103.9009
20.0212 | -103.9087
20.0249 | -103.9164
20.0286 | -103.9231
20.0323 | -103.9293
20.036 | -103.9346
20.0397 | -103.939
20.0433 | -103.9442
20.047 | -103.9475
20.0507 | -103.9513
20.0544 | -103.9551
20.0581 | -103.9599
20.0618 | -103.9637
20.0655 | -103.9669
20.0692 | -103.9714
20.0728 | -103.9729
20.0765 | -103.9727
20.0802 | -103.9751
20.0839 | -103.9781
20.0876 | -103.981
20.0913 | -103.9808
20.095 | -103.9817
20.0987 | -103.9816
20.1024 | -103.9819
20.106 | -103.9826
20.1097 | -103.9837
20.1134 | -103.984
20.1171 | -103.984
20.1208 | -103.9865
20.1245 | -103.9888
20.1282 | -103.99
20.1319 | -103.9862
20.1356 | -103.9822
20.1392 | -103.9791
20.1429 | -103.9759
20.1466 | -103.9725
20.1503 | -103.9698
20.154 | -103.9669
20.1577 | -103.9633
20.1614 | -103.961
20.1651 | -103.9575
20.1687 | -103.9527
20.1724 | -103.9491
20.1761 | -103.9445
20.1798 | -103.9403
20.1835 | -103.9364
20.1872 | -103.9323
20.1909 | -103.9291
20.1946 | -103.9266
20.1983 | -103.9222
20.2019 | -103.9176
20.2056 | -103.9147
20.2093 | -103.9111
20.213 | -103.9078
20.2167 | -103.9047
20.2204 | -103.9027
20.2241 | -103.9011
20.2278 | -103.8975
20.2315 | -103.8958
20.2351 | -103.8946
20.2388 | -103.8957
20.2425 | -103.8979
20.2462 | -103.8969
20.2499 | -103.8955
20.2536 | -103.8934
20.2573 | -103.8945
20.261 | -103.8972
20.2646 | -103.8971
20.2683 | -103.8961
20.272 | -103.8947
20.2757 | -103.8969
20.2794 | -103.9013
20.2831 | -103.9038
20.2868 | -103.9039
20.2905 | -103.9059
20.2942 | -103.9074
20.2978 | -103.9094
20.3015 | -103.9104
20.3052 | -103.9129
20.3089 | -103.9154
20.3126 | -103.9172
20.3163 | -103.9213
20.32 | -103.9249
20.3237 | -103.9278
20.3274 | -103.9309
20.331 | -103.9294
20.3347 | -103.9301
20.3384 | -103.9298
20.3421 | -103.927
20.3458 | -103.925
20.3495 | -103.9217
20.3532 | -103.9186
20.3569 | -103.916
20.3606 | -103.9146
20.3642 | -103.91
20.3679 | -103.9044
20.3716 | -103.8981
20.3753 | -103.891
20.379 | -103.8863
20.3827 | -103.8791
20.3864 | -103.871
20.3901 | -103.8627
20.3937 | -103.8544
20.3974 | -103.8455
20.4011 | -103.8347
20.4048 | -103.8237
20.4085 | -103.8112
20.4122 | -103.7981
20.4159 | -103.7877
20.4196 | -103.78
20.4233 | -103.7685
20.4269 | -103.7568
20.4306 | -103.7466
20.4343 | -103.7315
20.438 | -103.7178
20.4417 | -103.7053
20.4454 | -103.6936
20.4491 | -103.681
20.4528 | -103.67
20.4565 | -103.6605
20.4601 | -103.6505
20.4638 | -103.6409
20.4675 | -103.6312
20.4712 | -103.6228
20.4749 | -103.6151
20.4786 | -103.6046
20.4823 | -103.5933
20.486 | -103.5837
20.4896 | -103.5755
20.4933 | -103.565
20.497 | -103.5548
20.5007 | -103.5473
20.5044 | -103.5397
20.5081 | -103.5352
20.5118 | -103.5302
20.5155 | -103.5245
20.5192 | -103.5183
20.5228 | -103.5111
20.5265 | -103.506
20.5302 | -103.5018
20.5339 | -103.4976
20.5376 | -103.496
20.5413 | -103.4944
20.545 | -103.4908
20.5487 | -103.4845
20.5524 | -103.4805
20.556 | -103.479
20.5597 | -103.4793
20.5634 | -103.4812
20.5671 | -103.4804
20.5708 | -103.48
20.5745 | -103.4805
20.5782 | -103.4822
20.5819 | -103.4839
20.5856 | -103.4863
20.5892 | -103.4878
20.5929 | -103.4907
20.5966 | -103.4938
20.6003 | -103.4981
20.604 | -103.5026
20.6077 | -103.506
20.6114 | -103.5077
20.6151 | -103.5107
20.6187 | -103.5156
20.6224 | -103.5232
20.6261 | -103.5287
20.6298 | -103.5327
20.6335 | -103.5397
20.6372 | -103.5453
20.6409 | -103.5522
20.6446 | -103.5577
20.6483 | -103.5614
20.6519 | -103.5659
20.6556 | -103.5726
20.6593 | -103.5788
20.663 | -103.5841
20.6667 | -103.5924
20.6704 | -103.5992
20.6741 | -103.6059
20.6778 | -103.6122
20.6815 | -103.6173
20.6851 | -103.6209
20.6888 | -103.6253
20.6925 | -103.6321
20.6962 | -103.6387
20.6999 | -103.6462
20.7036 | -103.653
20.7073 | -103.6576
20.711 | -103.6637
20.7146 | -103.6721
20.7183 | -103.6777
20.722 | -103.6814
20.7257 | -103.6899
20.7294 | -103.6969
20.7331 | -103.7007
20.7368 | -103.7039
20.7405 | -103.7078
20.7442 | -103.7109
20.7478 | -103.7144
20.7515 | -103.7188
20.7552 | -103.7234
20.7589 | -103.7283
20.7626 | -103.7334
20.7663 | -103.7384
20.77 | -103.7436
20.7737 | -103.7492
20.7774 | -103.7549
20.781 | -103.7607
20.7847 | -103.7654
20.7884 | -103.7714
20.7921 | -103.7778
20.7958 | -103.7837
20.7995 | -103.7883
20.8032 | -103.7899
20.8069 | -103.7939
20.8105 | -103.7997
20.8142 | -103.8063
20.8179 | -103.8136
20.8216 | -103.8202
20.8253 | -103.8267
20.829 | -103.8341
20.8327 | -103.8391
20.8364 | -103.8439
20.8401 | -103.8522
20.8437 | -103.8628
20.8474 | -103.8695
20.8511 | -103.8776
20.8548 | -103.8852
20.8585 | -103.8913
20.8622 | -103.8989
20.8659 | -103.9076
20.8696 | -103.9176
20.8733 | -103.9284
20.8769 | -103.9381
20.8806 | -103.9462
20.8843 | -103.9561
20.888 | -103.964
20.8917 | -103.9698
20.8954 | -103.9766
20.8991 | -103.9847
20.9028 | -103.9941
20.9064 | -104.0043
20.9101 | -104.013
20.9138 | -104.0198
20.9175 | -104.0261
20.9212 | -104.0325
20.9249 | -104.0393
20.9286 | -104.0424
20.9323 | -104.0451
20.936 | -104.0475
20.9396 | -104.0511
20.9433 | -104.0563
20.947 | -104.0585
20.9507 | -104.0586
20.9544 | -104.0556
20.9581 | -104.0538
20.9618 | -104.0521
20.9655 | -104.0505
20.9692 | -104.048
20.9728 | -104.0458
20.9765 | -104.0427
20.9802 | -104.0409
20.9839 | -104.0374
20.9876 | -104.03
20.9913 | -104.0221
20.995 | -104.0155
20.9987 | -104.009
21.0024 | -104.0033
21.006 | -103.9955
21.0097 | -103.987
21.0134 | -103.9814
21.0171 | -103.9753
21.0208 | -103.9678
21.0245 | -103.9623
21.0282 | -103.9585
21.0319 | -103.9549
21.0355 | -103.9502
21.0392 | -103.9461
21.0429 | -103.9416
21.0466 | -103.9375
21.0503 | -103.9328
21.054 | -103.9305
21.0577 | -103.9277
21.0614 | -103.9246
21.0651 | -103.9219
21.0687 | -103.9218
21.0724 | -103.9235
21.0761 | -103.9258
21.0798 | -103.9281
21.0835 | -103.93
21.0872 | -103.9338
21.0909 | -103.9389
21.0946 | -103.9425
21.0983 | -103.948
21.1019 | -103.957
21.1056 | -103.9649
21.1093 | -103.9735
21.113 | -103.9851
21.1167 | -103.9955
21.1204 | -104.0061
21.1241 | -104.0185
21.1278 | -104.0318
21.1314 | -104.0478
21.1351 | -104.0647
21.1388 | -104.0824
21.1425 | -104.1001
21.1462 | -104.1188
21.1499 | -104.1375
21.1536 | -104.1564
21.1573 | -104.1752
21.161 | -104.1947
21.1646 | -104.2125
21.1683 | -104.2311
21.172 | -104.2503
21.1757 | -104.2702
21.1794 | -104.2893
21.1831 | -104.3056
21.1868 | -104.3255
21.1905 | -104.3457
21.1942 | -104.3618
21.1978 | -104.3767
21.2015 | -104.3929
21.2052 | -104.4078
21.2089 | -104.4214
21.2126 | -104.4352
21.2163 | -104.4486
21.22 | -104.4613
21.2237 | -104.4749
21.2274 | -104.4883
21.231 | -104.5
21.2347 | -104.5092
21.2384 | -104.5162
21.2421 | -104.5211
21.2458 | -104.5269
21.2495 | -104.535
21.2532 | -104.5436
21.2569 | -104.5475
21.2605 | -104.5488
21.2642 | -104.5518
21.2679 | -104.5537
21.2716 | -104.5547
21.2753 | -104.5577
21.279 | -104.5592
21.2827 | -104.557
21.2864 | -104.5577
21.2901 | -104.5614
21.2937 | -104.5642
21.2974 | -104.5646
21.3011 | -104.5634
21.3048 | -104.5623
21.3085 | -104.5625
21.3122 | -104.5637
21.3159 | -104.5628
21.3196 | -104.5604
21.3233 | -104.5605
21.3269 | -104.5604
21.3306 | -104.559
21.3343 | -104.5568
21.338 | -104.5555
21.3417 | -104.5559
21.3454 | -104.5577
21.3491 | -104.5576
21.3528 | -104.5568
21.3564 | -104.5537
21.3601 | -104.5532
21.3638 | -104.5541
21.3675 | -104.5541
21.3712 | -104.553
21.3749 | -104.5534
21.3786 | -104.5532
21.3823 | -104.5541
21.386 | -104.555
21.3896 | -104.5535
21.3933 | -104.5526
21.397 | -104.5532
21.4007 | -104.5529
21.4044 | -104.5516
21.4081 | -104.5521
21.4118 | -104.5527
21.4155 | -104.5505
21.4192 | -104.55
21.4228 | -104.5508
21.4265 | -104.5515
21.4302 | -104.5526
21.4339 | -104.5518
21.4376 | -104.5524
21.4413 | -104.5518
21.445 | -104.5516
21.4487 | -104.5509
21.4523 | -104.55
21.456 | -104.5497
21.4597 | -104.5498
21.4634 | -104.5508
21.4671 | -104.5538
21.4708 | -104.5564
21.4745 | -104.5572
21.4782 | -104.5587
21.4819 | -104.5622
21.4855 | -104.5662
21.4892 | -104.568
21.4929 | -104.5692
21.4966 | -104.5723
21.5003 | -104.5748
21.504 | -104.5748
21.5077 | -104.576
21.5114 | -104.5757
21.5151 | -104.5753
21.5187 | -104.5738
21.5224 | -104.572
21.5261 | -104.5701
21.5298 | -104.568
21.5335 | -104.565
21.5372 | -104.5634
21.5409 | -104.5608
21.5446 | -104.5585
21.5482 | -104.5536
21.5519 | -104.5471
21.5556 | -104.5415
21.5593 | -104.5346
21.563 | -104.5281
21.5667 | -104.5208
21.5704 | -104.5107
21.5741 | -104.5007
21.5778 | -104.49
21.5814 | -104.476
21.5851 | -104.4621
21.5888 | -104.4481
21.5925 | -104.4326
21.5962 | -104.4155
21.5999 | -104.3989
21.6036 | -104.3831
21.6073 | -104.3689
21.611 | -104.3537
21.6146 | -104.3364
21.6183 | -104.3185
21.622 | -104.3024
21.6257 | -104.2869
21.6294 | -104.2672
21.6331 | -104.2488
21.6368 | -104.2317
21.6405 | -104.2151
21.6442 | -104.1959
21.6478 | -104.1762
21.6515 | -104.1565
21.6552 | -104.1399
21.6589 | -104.1255
21.6626 | -104.1105
21.6663 | -104.094
21.67 | -104.0757
21.6737 | -104.0579
21.6773 | -104.0412
21.681 | -104.0261
21.6847 | -104.0119
21.6884 | -103.9962
21.6921 | -103.9824
21.6958 | -103.9685
21.6995 | -103.9548
21.7032 | -103.9435
21.7069 | -103.9328
21.7105 | -103.9215
21.7142 | -103.9097
21.7179 | -103.8974
21.7216 | -103.8858
21.7253 | -103.8755
21.729 | -103.8645
21.7327 | -103.8517
21.7364 | -103.8403
21.7401 | -103.8322
21.7437 | -103.8242
21.7474 | -103.8154
21.7511 | -103.8071
21.7548 | -103.7984
21.7585 | -103.7896
21.7622 | -103.7827
21.7659 | -103.7767
21.7696 | -103.7701
21.7732 | -103.7617
21.7769 | -103.7534
21.7806 | -103.7465
21.7843 | -103.7393
21.788 | -103.7315
21.7917 | -103.7244
21.7954 | -103.7192
21.7991 | -103.7148
21.8028 | -103.7092
21.8064 | -103.7032
21.8101 | -103.6961
21.8138 | -103.6883
21.8175 | -103.6834
21.8212 | -103.6801
21.8249 | -103.6751
21.8286 | -103.6701
21.8323 | -103.6653
21.836 | -103.6603
21.8396 | -103.6565
21.8433 | -103.6531
21.847 | -103.6465
21.8507 | -103.6418
21.8544 | -103.6359
21.8581 | -103.6302
21.8618 | -103.6258
21.8655 | -103.6197
21.8692 | -103.6154
21.8728 | -103.6112
21.8765 | -103.6048
21.8802 | -103.598
21.8839 | -103.5901
21.8876 | -103.5828
21.8913 | -103.5787
21.895 | -103.573
21.8987 | -103.564
21.9023 | -103.557
21.906 | -103.55
21.9097 | -103.5411
21.9134 | -103.5307
21.9171 | -103.5211
21.9208 | -103.5128
21.9245 | -103.5046
21.9282 | -103.4963
21.9319 | -103.4871
21.9355 | -103.4798
21.9392 | -103.4734
21.9429 | -103.4643
21.9466 | -103.4559
21.9503 | -103.4478
21.954 | -103.4414
21.9577 | -103.4349
21.9614 | -103.4296
21.9651 | -103.4233
21.9687 | -103.4154
21.9724 | -103.4085
21.9761 | -103.4032
21.9798 | -103.3993
21.9835 | -103.3962
21.9872 | -103.392
21.9909 | -103.3849
21.9946 | -103.3793
21.9982 | -103.3767
22.0019 | -103.3725
22.0056 | -103.3707
22.0093 | -103.3699
22.013 | -103.3677
22.0167 | -103.3667
22.0204 | -103.3662
22.0241 | -103.3657
22.0278 | -103.3658
22.0314 | -103.3677
22.0351 | -103.3698
22.0388 | -103.3739
22.0425 | -103.3792
22.0462 | -103.3839
22.0499 | -103.3871
22.0536 | -103.3905
22.0573 | -103.3946
22.061 | -103.399
22.0646 | -103.404
22.0683 | -103.4096
22.072 | -103.4111
22.0757 | -103.4094
22.0794 | -103.4095
22.0831 | -103.4131
22.0868 | -103.4157
22.0905 | -103.4175
22.0941 | -103.4183
22.0978 | -103.4196
22.1015 | -103.417
22.1052 | -103.414
22.1089 | -103.4125
22.1126 | -103.41
22.1163 | -103.4054
22.12 | -103.4011
22.1237 | -103.3966
22.1273 | -103.3918
22.131 | -103.3855
22.1347 | -103.3783
22.1384 | -103.369
22.1421 | -103.3594
22.1458 | -103.3509
22.1495 | -103.3412
22.1532 | -103.3301
22.1569 | -103.3179
22.1605 | -103.3065
22.1642 | -103.2949
22.1679 | -103.2832
22.1716 | -103.2713
22.1753 | -103.2585
22.179 | -103.2478
22.1827 | -103.2349
22.1864 | -103.221
22.19 | -103.2089
22.1937 | -103.2006
22.1974 | -103.1913
22.2011 | -103.181
22.2048 | -103.1692
22.2085 | -103.16
22.2122 | -103.1538
22.2159 | -103.1452
22.2196 | -103.134
22.2232 | -103.1247
22.2269 | -103.1189
22.2306 | -103.1124
22.2343 | -103.1047
22.238 | -103.0986
22.2417 | -103.094
22.2454 | -103.0903
22.2491 | -103.0856
22.2528 | -103.082
22.2564 | -103.0801
22.2601 | -103.0782
22.2638 | -103.0774
22.2675 | -103.0765
22.2712 | -103.0735
22.2749 | -103.0683
22.2786 | -103.0669
22.2823 | -103.0667
22.286 | -103.0661
22.2896 | -103.0641
22.2933 | -103.0606
22.297 | -103.0559
22.3007 | -103.0516
22.3044 | -103.0496
22.3081 | -103.0472
22.3118 | -103.0431
22.3155 | -103.0384
22.3191 | -103.0327
22.3228 | -103.0267
22.3265 | -103.0201
22.3302 | -103.0142
22.3339 | -103.0083
22.3376 | -103.003
22.3413 | -102.9977
22.345 | -102.9902
22.3487 | -102.9792
22.3523 | -102.97
22.356 | -102.961
22.3597 | -102.9518
22.3634 | -102.9413
22.3671 | -102.9334
22.3708 | -102.926
22.3745 | -102.9193
22.3782 | -102.9108
22.3819 | -102.9029
22.3855 | -102.8957
22.3892 | -102.8905
22.3929 | -102.8862
22.3966 | -102.8812
22.4003 | -102.8737
22.404 | -102.8659
22.4077 | -102.8598
22.4114 | -102.8542
22.415 | -102.8491
22.4187 | -102.8473
22.4224 | -102.8453
22.4261 | -102.8402
22.4298 | -102.8378
22.4335 | -102.8381
22.4372 | -102.8384
22.4409 | -102.8395
22.4446 | -102.8389
22.4482 | -102.8372
22.4519 | -102.8371
22.4556 | -102.836
22.4593 | -102.8319
22.463 | -102.8256
22.4667 | -102.8226
22.4704 | -102.8176
22.4741 | -102.8143
22.4778 | -102.8075
22.4814 | -102.7952
22.4851 | -102.7773
22.4888 | -102.7566
22.4925 | -102.7354
22.4962 | -102.711
22.4999 | -102.6901
22.5036 | -102.6693
22.5073 | -102.6465
22.511 | -102.6179
22.5146 | -102.5894
22.5183 | -102.5617
22.522 | -102.532
22.5257 | -102.4982
22.5294 | -102.4604
22.5331 | -102.4235
22.5368 | -102.3871
22.5405 | -102.3471
22.5441 | -102.3041
22.5478 | -102.2589
22.5515 | -102.2144
22.5552 | -102.1677
22.5589 | -102.1209
22.5626 | -102.0736
22.5663 | -102.0262
22.57 | -101.978
22.5737 | -101.9302
22.5773 | -101.8816
22.581 | -101.835
22.5847 | -101.7903
22.5884 | -101.7467
22.5921 | -101.7052
22.5958 | -101.6628
22.5995 | -101.6218
22.6032 | -101.582
22.6069 | -101.5468
22.6105 | -101.5153
22.6142 | -101.4851
22.6179 | -101.4575
22.6216 | -101.4321
22.6253 | -101.4098
22.629 | -101.3896
22.6327 | -101.3729
22.6364 | -101.3599
22.64 | -101.3496
22.6437 | -101.3429
22.6474 | -101.34
22.6511 | -101.3389
22.6548 | -101.3415
22.6585 | -101.3486
22.6622 | -101.3587
22.6659 | -101.3717
22.6696 | -101.3887
22.6732 | -101.4086
22.6769 | -101.429
22.6806 | -101.4526
22.6843 | -101.4803
22.688 | -101.5097
22.6917 | -101.5401
22.6954 | -101.5722
22.6991 | -101.6076
22.7028 | -101.6439
22.7064 | -101.6791
22.7101 | -101.7159
22.7138 | -101.7549
22.7175 | -101.7933
22.7212 | -101.833
22.7249 | -101.873
22.7286 | -101.9104
22.7323 | -101.9457
22.7359 | -101.9799
22.7396 | -102.0128
22.7433 | -102.0452
22.747 | -102.0766
22.7507 | -102.105
22.7544 | -102.1319
22.7581 | -102.1571
22.7618 | -102.1802
22.7655 | -102.1989
22.7691 | -102.2166
22.7728 | -102.2334
22.7765 | -102.2502
22.7802 | -102.2664
22.7839 | -102.2786
22.7876 | -102.2862
22.7913 | -102.292
22.795 | -102.2939
22.7987 | -102.2984
22.8023 | -102.2989
22.806 | -102.3005
22.8097 | -102.3027
22.8134 | -102.2985
22.8171 | -102.2937
22.8208 | -102.2831
22.8245 | -102.2676
22.8282 | -102.2515
22.8318 | -102.2363
22.8355 | -102.217
22.8392 | -102.1909
22.8429 | -102.1673
22.8466 | -102.147
22.8503 | -102.1272
22.854 | -102.1036
22.8577 | -102.0807
22.8614 | -102.059
22.865 | -102.0359
22.8687 | -102.0166
22.8724 | -101.9959
22.8761 | -101.9749
22.8798 | -101.96
22.8835 | -101.9395
22.8872 | -101.9208
22.8909 | -101.902
22.8946 | -101.8828
22.8982 | -101.8649
22.9019 | -101.8482
22.9056 | -101.8287
22.9093 | -101.8093
22.913 | -101.7913
22.9167 | -101.7737
22.9204 | -101.7582
22.9241 | -101.7444
22.9278 | -101.7334
22.9314 | -101.7224
22.9351 | -101.7115
22.9388 | -101.7019
22.9425 | -101.6929
22.9462 | -101.6855
22.9499 | -101.6773
22.9536 | -101.6708
22.9573 | -101.6669
22.9609 | -101.6618
22.9646 | -101.6562
22.9683 | -101.6501
22.972 | -101.6449
22.9757 | -101.6426
22.9794 | -101.6381
22.9831 | -101.6324
22.9868 | -101.6292
22.9905 | -101.6251
22.9941 | -101.6184
22.9978 | -101.6116
23.0015 | -101.6062
23.0052 | -101.6008
23.0089 | -101.5962
23.0126 | -101.5922
23.0163 | -101.5902
23.02 | -101.5867
23.0237 | -101.5806
23.0273 | -101.5746
23.031 | -101.5697
23.0347 | -101.5639
23.0384 | -101.5569
23.0421 | -101.5505
23.0458 | -101.5455
23.0495 | -101.5403
23.0532 | -101.5346
23.0568 | -101.5296
23.0605 | -101.5224
23.0642 | -101.5148
23.0679 | -101.507
23.0716 | -101.5006
23.0753 | -101.4927
23.079 | -101.4829
23.0827 | -101.474
23.0864 | -101.4633
23.09 | -101.4546
23.0937 | -101.4459
23.0974 | -101.435
23.1011 | -101.4252
23.1048 | -101.4158
23.1085 | -101.4066
23.1122 | -101.3965
23.1159 | -101.3878
23.1196 | -101.3788
23.1232 | -101.367
23.1269 | -101.3584
23.1306 | -101.35
23.1343 | -101.3395
23.138 | -101.3324
23.1417 | -101.3254
23.1454 | -101.3184
23.1491 | -101.3139
23.1528 | -101.3085
23.1564 | -101.3054
23.1601 | -101.3036
23.1638 | -101.3061
23.1675 | -101.3083
23.1712 | -101.3078
23.1749 | -101.3084
23.1786 | -101.312
23.1823 | -101.3166
23.1859 | -101.3225
23.1896 | -101.3292
23.1933 | -101.3338
23.197 | -101.3405
23.2007 | -101.3504
23.2044 | -101.361
23.2081 | -101.3717
23.2118 | -101.3828
23.2155 | -101.3909
23.2191 | -101.3988
23.2228 | -101.4047
23.2265 | -101.4113
23.2302 | -101.4201
23.2339 | -101.4302
23.2376 | -101.4401
23.2413 | -101.4516
23.245 | -101.4616
23.2487 | -101.4694
23.2523 | -101.476
23.256 | -101.4823
23.2597 | -101.4858
23.2634 | -101.4876
23.2671 | -101.49
23.2708 | -101.4914
23.2745 | -101.4913
23.2782 | -101.4922
23.2818 | -101.4925
23.2855 | -101.4919
23.2892 | -101.4906
23.2929 | -101.4889
23.2966 | -101.4861
23.3003 | -101.479
23.304 | -101.4709
23.3077 | -101.4651
23.3114 | -101.4595
23.315 | -101.4528
23.3187 | -101.447
23.3224 | -101.4416
23.3261 | -101.4328
23.3298 | -101.4225
23.3335 | -101.4121
23.3372 | -101.3994
23.3409 | -101.3851
23.3446 | -101.3704
23.3482 | -101.3564
23.3519 | -101.3463
23.3556 | -101.3372
23.3593 | -101.3274
23.363 | -101.3179
23.3667 | -101.3105
23.3704 | -101.3032
23.3741 | -101.2945
23.3777 | -101.2876
23.3814 | -101.2824
23.3851 | -101.2784
23.3888 | -101.2767
23.3925 | -101.2746
23.3962 | -101.2715
23.3999 | -101.2679
23.4036 | -101.2683
23.4073 | -101.2689
23.4109 | -101.276
23.4146 | -101.2785
23.4183 | -101.2837
23.422 | -101.2893
23.4257 | -101.294
23.4294 | -101.2995
23.4331 | -101.305
23.4368 | -101.3109
23.4405 | -101.3167
23.4441 | -101.3213
23.4478 | -101.3237
23.4515 | -101.3267
23.4552 | -101.3349
23.4589 | -101.3432
23.4626 | -101.351
23.4663 | -101.3576
23.47 | -101.3614
23.4736 | -101.3634
23.4773 | -101.3664
23.481 | -101.3694
23.4847 | -101.3732
23.4884 | -101.3757
23.4921 | -101.3782
23.4958 | -101.3786
23.4995 | -101.3764
23.5032 | -101.3709
23.5068 | -101.3665
23.5105 | -101.361
23.5142 | -101.3536
23.5179 | -101.3449
23.5216 | -101.3369
23.5253 | -101.3292
23.529 | -101.3218
23.5327 | -101.3124
23.5364 | -101.302
23.54 | -101.2925
23.5437 | -101.2816
23.5474 | -101.2712
23.5511 | -101.2621
23.5548 | -101.2527
23.5585 | -101.2408
23.5622 | -101.2263
23.5659 | -101.2117
23.5696 | -101.197
23.5732 | -101.1808
23.5769 | -101.1661
23.5806 | -101.1535
23.5843 | -101.1413
23.588 | -101.1272
23.5917 | -101.1134
23.5954 | -101.1011
23.5991 | -101.0883
23.6027 | -101.076
23.6064 | -101.0656
23.6101 | -101.0567
23.6138 | -101.0491
23.6175 | -101.0397
23.6212 | -101.0311
23.6249 | -101.0259
23.6286 | -101.023
23.6323 | -101.0209
23.6359 | -101.0202
23.6396 | -101.0198
23.6433 | -101.0203
23.647 | -101.02
23.6507 | -101.0208
23.6544 | -101.0248
23.6581 | -101.0319
23.6618 | -101.0407
23.6655 | -101.0478
23.6691 | -101.0547
23.6728 | -101.0638
23.6765 | -101.0733
23.6802 | -101.0851
23.6839 | -101.0955
23.6876 | -101.1058
23.6913 | -101.1174
23.695 | -101.1299
23.6986 | -101.1422
23.7023 | -101.1546
23.706 | -101.1683
23.7097 | -101.1833
23.7134 | -101.1984
23.7171 | -101.2134
23.7208 | -101.228
23.7245 | -101.2424
23.7282 | -101.2574
23.7318 | -101.2731
23.7355 | -101.2869
23.7392 | -101.3028
23.7429 | -101.3189
23.7466 | -101.3346
23.7503 | -101.349
23.754 | -101.3626
23.7577 | -101.3765
23.7614 | -101.3913
23.765 | -101.4062
23.7687 | -101.421
23.7724 | -101.4355
23.7761 | -101.4528
23.7798 | -101.4712
23.7835 | -101.4913
23.7872 | -101.5108
23.7909 | -101.5248
23.7946 | -101.5387
23.7982 | -101.5534
23.8019 | -101.5697
23.8056 | -101.5865
23.8093 | -101.6052
23.813 | -101.6243
23.8167 | -101.644
23.8204 | -101.6625
23.8241 | -101.6809
23.8277 | -101.7
23.8314 | -101.72
23.8351 | -101.7427
23.8388 | -101.7689
23.8425 | -101.7934
23.8462 | -101.8174
23.8499 | -101.8398
23.8536 | -101.8633
23.8573 | -101.8859
23.8609 | -101.9071
23.8646 | -101.9289
23.8683 | -101.9516
23.872 | -101.9741
23.8757 | -101.9979
23.8794 | -102.0248
23.8831 | -102.0508
23.8868 | -102.0751
23.8905 | -102.1025
23.8941 | -102.1285
23.8978 | -102.1544
23.9015 | -102.1802
23.9052 | -102.2043
23.9089 | -102.2283
23.9126 | -102.2507
23.9163 | -102.2707
23.92 | -102.2924
23.9236 | -102.316
23.9273 | -102.3409
23.931 | -102.3649
23.9347 | -102.3906
23.9384 | -102.4157
23.9421 | -102.4378
23.9458 | -102.4525
23.9495 | -102.4709
23.9532 | -102.49
23.9568 | -102.5121
23.9605 | -102.5349
23.9642 | -102.555
23.9679 | -102.5747
23.9716 | -102.5948
23.9753 | -102.6104
23.979 | -102.6243
23.9827 | -102.6399
23.9864 | -102.6567
23.99 | -102.6682
23.9937 | -102.6771
23.9974 | -102.688
24.0011 | -102.6983
24.0048 | -102.7104
24.0085 | -102.7219
24.0122 | -102.7318
24.0159 | -102.7399
24.0195 | -102.7452
24.0232 | -102.7498
24.0269 | -102.7518
24.0306 | -102.7518
24.0343 | -102.7526
24.038 | -102.7524
24.0417 | -102.7524
24.0454 | -102.7511
24.0491 | -102.7477
24.0527 | -102.7441
24.0564 | -102.7407
24.0601 | -102.7372
24.0638 | -102.7332
24.0675 | -102.7301
24.0712 | -102.722
24.0749 | -102.7085
24.0786 | -102.6976
24.0823 | -102.6889
24.0859 | -102.6783
24.0896 | -102.6656
24.0933 | -102.6537
24.097 | -102.641
24.1007 | -102.6275
24.1044 | -102.6135
24.1081 | -102.5975
24.1118 | -102.5831
24.1154 | -102.5683
24.1191 | -102.555
24.1228 | -102.5402
24.1265 | -102.5255
24.1302 | -102.5118
24.1339 | -102.4981
24.1376 | -102.4842
24.1413 | -102.4708
24.145 | -102.4562
24.1486 | -102.4411
24.1523 | -102.4255
24.156 | -102.4099
24.1597 | -102.3959
24.1634 | -102.3821
24.1671 | -102.369
24.1708 | -102.3542
24.1745 | -102.3392
24.1782 | -102.325
24.1818 | -102.3118
24.1855 | -102.2984
24.1892 | -102.2844
24.1929 | -102.2711
24.1966 | -102.2599
24.2003 | -102.2509
24.204 | -102.2396
24.2077 | -102.2261
24.2114 | -102.2142
24.215 | -102.2001
24.2187 | -102.1879
24.2224 | -102.1792
24.2261 | -102.1702
24.2298 | -102.1546
24.2335 | -102.1378
24.2372 | -102.1223
24.2409 | -102.1054
24.2445 | -102.0798
24.2482 | -102.0422
24.2519 | -101.9914
24.2556 | -101.9347
24.2593 | -101.8655
24.263 | -101.7944
24.2667 | -101.7316
24.2704 | -101.6748
24.2741 | -101.6199
24.2777 | -101.5631
24.2814 | -101.5046
24.2851 | -101.4429
24.2888 | -101.3783
24.2925 | -101.3116
24.2962 | -101.2416
24.2999 | -101.1674
24.3036 | -101.0911
24.3073 | -101.0132
24.3109 | -100.9315
24.3146 | -100.8471
24.3183 | -100.7642
24.322 | -100.6779
24.3257 | -100.5883
24.3294 | -100.4995
24.3331 | -100.4122
24.3368 | -100.3242
24.3404 | -100.2352
24.3441 | -100.1458
24.3478 | -100.0563
24.3515 | -99.9689
24.3552 | -99.8826
24.3589 | -99.7984
24.3626 | -99.7168
24.3663 | -99.6371
24.37 | -99.5617
24.3736 | -99.4902
24.3773 | -99.4213
24.381 | -99.3546
24.3847 | -99.2919
24.3884 | -99.2343
24.3921 | -99.1808
24.3958 | -99.132
24.3995 | -99.0872
24.4032 | -99.0473
24.4068 | -99.0135
24.4105 | -98.9843
24.4142 | -98.9615
24.4179 | -98.9461
24.4216 | -98.9329
24.4253 | -98.9244
24.429 | -98.922
24.4327 | -98.9255
24.4364 | -98.9346
24.44 | -98.9495
24.4437 | -98.9701
24.4474 | -98.9958
24.4511 | -99.0264
24.4548 | -99.0618
24.4585 | -99.1014
24.4622 | -99.1455
24.4659 | -99.1941
24.4695 | -99.2469
24.4732 | -99.3048
24.4769 | -99.3893
24.4806 | -99.4515
24.4843 | -99.5171
24.488 | -99.5853
24.4917 | -99.6558
24.4954 | -99.7293
24.4991 | -99.8054
24.5027 | -99.8832
24.5064 | -99.9638
24.5101 | -100.0458
24.5138 | -100.1286
24.5175 | -100.2132
24.5212 | -100.2972
24.5249 | -100.3817
24.5286 | -100.4662
24.5323 | -100.5507
24.5359 | -100.6355
24.5396 | -100.7181
24.5433 | -100.7988
24.547 | -100.8777
24.5507 | -100.9559
24.5544 | -101.0338
24.5581 | -101.1095
24.5618 | -101.1823
24.5654 | -101.2495
24.5691 | -101.3084
24.5728 | -101.3584
24.5765 | -101.3937
24.5802 | -101.3699
24.5839 | -101.3262
24.5876 | -101.2703
24.5913 | -101.1809
24.595 | -101.1263
24.5986 | -101.0988
24.6023 | -101.0758
24.606 | -101.0487
24.6097 | -101.0076
24.6134 | -100.9452
24.6171 | -100.8654
24.6208 | -100.772
24.6245 | -100.6714
24.6282 | -100.5652
24.6318 | -100.4497
24.6355 | -100.3289
24.6392 | -100.2044
24.6429 | -100.0784
24.6466 | -99.9505
24.6503 | -99.8189
24.654 | -99.6849
24.6577 | -99.5494
24.6613 | -99.4141
24.665 | -99.2764
24.6687 | -99.1407
24.6724 | -99.005
24.6761 | -98.8695
24.6798 | -98.736
24.6835 | -98.6041
24.6872 | -98.4744
24.6909 | -98.3483
24.6945 | -98.2265
24.6982 | -98.1096
24.7019 | -97.9968
24.7056 | -97.8883
24.7093 | -97.7852
24.713 | -97.6883
24.7167 | -97.5969
24.7204 | -97.5111
24.7241 | -97.431
24.7277 | -97.3586
24.7314 | -97.2927
24.7351 | -97.2327
24.7388 | -97.1806
24.7425 | -97.1357
24.7462 | -97.0984
24.7499 | -97.0683
24.7536 | -97.0452
24.7572 | -97.0296
24.7609 | -97.0218
24.7646 | -97.0212
24.7683 | -97.0285
24.772 | -97.0432
24.7757 | -97.0654
24.7794 | -97.0953
24.7831 | -97.1325
24.7868 | -97.1771
24.7904 | -97.2286
24.7941 | -97.2876
24.7978 | -97.3542
24.8015 | -97.4277
24.8052 | -97.5079
24.8089 | -97.5947
24.8126 | -97.6881
24.8163 | -97.7876
24.82 | -97.8929
24.8236 | -98.0035
24.8273 | -98.1205
24.831 | -98.2429
24.8347 | -98.3697
24.8384 | -98.5016
24.8421 | -98.6389
24.8458 | -98.7795
24.8495 | -98.9232
24.8532 | -99.0699
24.8568 | -99.2189
24.8605 | -99.3704
24.8642 | -99.5231
24.8679 | -99.6763
24.8716 | -99.8312
24.8753 | -99.9857
24.879 | -100.1389
24.8827 | -100.2921
24.8863 | -100.4437
24.89 | -100.5933
24.8937 | -100.74
24.8974 | -100.8831
24.9011 | -101.0218
24.9048 | -101.1563
24.9085 | -101.2855
24.9122 | -101.4136
24.9159 | -101.5386
24.9195 | -101.6615
24.9232 | -101.7834
24.9269 | -101.9121
24.9306 | -102.0876
24.9343 | -102.2574
24.938 | -102.4121
24.9417 | -102.5854
24.9454 | -102.6851
24.9491 | -102.7306
24.9527 | -102.7515
24.9564 | -102.7657
24.9601 | -102.7783
24.9638 | -102.7905
24.9675 | -102.8016
24.9712 | -102.812
24.9749 | -102.819
24.9786 | -102.8231
24.9822 | -102.8316
24.9859 | -102.8389
24.9896 | -102.841
24.9933 | -102.8385
24.997 | -102.8326
25.0007 | -102.828
25.0044 | -102.8231
25.0081 | -102.8189
25.0118 | -102.815
25.0154 | -102.8106
25.0191 | -102.8049
25.0228 | -102.7977
25.0265 | -102.7921
25.0302 | -102.7854
25.0339 | -102.7797
25.0376 | -102.7768
25.0413 | -102.7731
25.045 | -102.7691
25.0486 | -102.7647
25.0523 | -102.7585
25.056 | -102.7523
25.0597 | -102.7487
25.0634 | -102.7464
25.0671 | -102.7406
25.0708 | -102.7359
25.0745 | -102.7332
25.0782 | -102.732
25.0818 | -102.7298
25.0855 | -102.7274
25.0892 | -102.7258
25.0929 | -102.7239
25.0966 | -102.7238
25.1003 | -102.7241
25.104 | -102.7237
25.1077 | -102.7233
25.1113 | -102.7237
25.115 | -102.7247
25.1187 | -102.7274
25.1224 | -102.7286
25.1261 | -102.7267
25.1298 | -102.7274
25.1335 | -102.7297
25.1372 | -102.7298
25.1409 | -102.7297
25.1445 | -102.7317
25.1482 | -102.7327
25.1519 | -102.7341
25.1556 | -102.7329
25.1593 | -102.7323
25.163 | -102.7342
25.1667 | -102.7363
25.1704 | -102.7371
25.1741 | -102.7359
25.1777 | -102.7354
25.1814 | -102.7323
25.1851 | -102.7267
25.1888 | -102.7224
25.1925 | -102.7187
25.1962 | -102.7126
25.1999 | -102.7058
25.2036 | -102.6987
25.2072 | -102.6917
25.2109 | -102.6844
25.2146 | -102.6771
25.2183 | -102.6709
25.222 | -102.6635
25.2257 | -102.6564
25.2294 | -102.6516
25.2331 | -102.6465
25.2368 | -102.6383
25.2404 | -102.6306
25.2441 | -102.6223
25.2478 | -102.6156
25.2515 | -102.6087
25.2552 | -102.6004
25.2589 | -102.5942
25.2626 | -102.5858
25.2663 | -102.5732
25.27 | -102.5616
25.2736 | -102.5503
25.2773 | -102.5389
25.281 | -102.5284
25.2847 | -102.5186
25.2884 | -102.509
25.2921 | -102.499
25.2958 | -102.4893
25.2995 | -102.4807
25.3031 | -102.4686
25.3068 | -102.4553
25.3105 | -102.4409
25.3142 | -102.4264
25.3179 | -102.412
25.3216 | -102.3979
25.3253 | -102.3846
25.329 | -102.3699
25.3327 | -102.356
25.3363 | -102.3413
25.34 | -102.3247
25.3437 | -102.3076
25.3474 | -102.2896
25.3511 | -102.2721
25.3548 | -102.2517
25.3585 | -102.2316
25.3622 | -102.2116
25.3659 | -102.1851
25.3695 | -102.1594
25.3732 | -102.1346
25.3769 | -102.1124
25.3806 | -102.0903
25.3843 | -102.0666
25.388 | -102.0416
25.3917 | -102.0166
25.3954 | -101.9897
25.399 | -101.9609
25.4027 | -101.9315
25.4064 | -101.9046
25.4101 | -101.8757
25.4138 | -101.8456
25.4175 | -101.8145
25.4212 | -101.783
25.4249 | -101.754
25.4286 | -101.7228
25.4322 | -101.6911
25.4359 | -101.659
25.4396 | -101.6271
25.4433 | -101.5927
25.447 | -101.5583
25.4507 | -101.5252
25.4544 | -101.4917
25.4581 | -101.4578
25.4618 | -101.4228
25.4654 | -101.389
25.4691 | -101.3573
25.4728 | -101.3233
25.4765 | -101.2894
25.4802 | -101.2569
25.4839 | -101.2221
25.4876 | -101.186
25.4913 | -101.1503
25.495 | -101.1171
25.4986 | -101.0835
25.5023 | -101.0422
25.506 | -101.0046
25.5097 | -100.9704
25.5134 | -100.9385
25.5171 | -100.9065
25.5208 | -100.8728
25.5245 | -100.8392
25.5281 | -100.8082
25.5318 | -100.7788
25.5355 | -100.7501
25.5392 | -100.7262
25.5429 | -100.6992
25.5466 | -100.6739
25.5503 | -100.6487
25.554 | -100.623
25.5577 | -100.6
25.5613 | -100.5782
25.565 | -100.5575
25.5687 | -100.5378
25.5724 | -100.5187
25.5761 | -100.5004
25.5798 | -100.4829
25.5835 | -100.4676
25.5872 | -100.4535
25.5909 | -100.4395
25.5945 | -100.4272
25.5982 | -100.4178
25.6019 | -100.4094
25.6056 | -100.4021
25.6093 | -100.3938
25.613 | -100.3875
25.6167 | -100.3836
25.6204 | -100.3803
25.624 | -100.3805
25.6277 | -100.3829
25.6314 | -100.3848
25.6351 | -100.3878
25.6388 | -100.3928
25.6425 | -100.3983
25.6462 | -100.4047
25.6499 | -100.4121
25.6536 | -100.4226
25.6572 | -100.4341
25.6609 | -100.4465
25.6646 | -100.4612
25.6683 | -100.4764
25.672 | -100.4929
25.6757 | -100.5104
25.6794 | -100.5284
25.6831 | -100.5473
25.6868 | -100.567
25.6904 | -100.589
25.6941 | -100.6085
25.6978 | -100.63
25.7015 | -100.6531
25.7052 | -100.6774
25.7089 | -100.7027
25.7126 | -100.7294
25.7163 | -100.7569
25.72 | -100.7833
25.7236 | -100.8094
25.7273 | -100.8404
25.731 | -100.8697
25.7347 | -100.8963
25.7384 | -100.9233
25.7421 | -100.95
25.7458 | -100.9768
25.7495 | -101.0046
25.7531 | -101.0325
25.7568 | -101.0604
25.7605 | -101.0892
25.7642 | -101.117
25.7679 | -101.1426
25.7716 | -101.1671
25.7753 | -101.1935
25.779 | -101.2196
25.7827 | -101.2447
25.7863 | -101.268
25.79 | -101.2906
25.7937 | -101.3112
25.7974 | -101.3325
25.8011 | -101.3526
25.8048 | -101.3751
25.8085 | -101.3967
25.8122 | -101.4146
25.8159 | -101.4296
25.8195 | -101.4458
25.8232 | -101.4633
25.8269 | -101.4784
25.8306 | -101.4902
25.8343 | -101.5028
25.838 | -101.5148
25.8417 | -101.5262
25.8454 | -101.5405
25.849 | -101.5522
25.8527 | -101.5622
25.8564 | -101.5717
25.8601 | -101.5804
25.8638 | -101.5964
25.8675 | -101.6067
25.8712 | -101.6125
25.8749 | -101.6188
25.8786 | -101.6261
25.8822 | -101.6339
25.8859 | -101.6412
25.8896 | -101.6474
25.8933 | -101.6524
25.897 | -101.6572
25.9007 | -101.6633
25.9044 | -101.6684
25.9081 | -101.6746
25.9118 | -101.682
25.9154 | -101.6888
25.9191 | -101.6948
25.9228 | -101.7008
25.9265 | -101.7074
25.9302 | -101.7146
25.9339 | -101.7203
25.9376 | -101.7244
25.9413 | -101.7293
25.9449 | -101.7351
25.9486 | -101.7389
25.9523 | -101.7353
25.956 | -101.7321
25.9597 | -101.7293
25.9634 | -101.7241
25.9671 | -101.7214
25.9708 | -101.7211
25.9745 | -101.7185
25.9781 | -101.7105
25.9818 | -101.7016
25.9855 | -101.6936
25.9892 | -101.6839
25.9929 | -101.6727
25.9966 | -101.6625
26.0003 | -101.6523
26.004 | -101.639
26.0077 | -101.6255
26.0113 | -101.61
26.015 | -101.5945
26.0187 | -101.5762
26.0224 | -101.5544
26.0261 | -101.5333
26.0298 | -101.5117
26.0335 | -101.4895
26.0372 | -101.4642
26.0408 | -101.438
26.0445 | -101.4107
26.0482 | -101.3846
26.0519 | -101.3579
26.0556 | -101.3307
26.0593 | -101.3058
26.063 | -101.2797
26.0667 | -101.2506
26.0704 | -101.2193
26.074 | -101.1859
26.0777 | -101.155
26.0814 | -101.125
26.0851 | -101.0978
26.0888 | -101.071
26.0925 | -101.045
26.0962 | -101.0202
26.0999 | -100.9956
26.1036 | -100.9735
26.1072 | -100.9498
26.1109 | -100.9263
26.1146 | -100.9037
26.1183 | -100.8836
26.122 | -100.8628
26.1257 | -100.84
26.1294 | -100.8196
26.1331 | -100.8017
26.1368 | -100.7844
26.1404 | -100.7706
26.1441 | -100.7565
26.1478 | -100.743
26.1515 | -100.7302
26.1552 | -100.7187
26.1589 | -100.7111
26.1626 | -100.7066
26.1663 | -100.7015
26.1699 | -100.6962
26.1736 | -100.691
26.1773 | -100.6862
26.181 | -100.6822
26.1847 | -100.6807
26.1884 | -100.6802
26.1921 | -100.6794
26.1958 | -100.6804
26.1995 | -100.6838
26.2031 | -100.6883
26.2068 | -100.6921
26.2105 | -100.695
26.2142 | -100.701
26.2179 | -100.7075
26.2216 | -100.7148
26.2253 | -100.7202
26.229 | -100.7272
26.2327 | -100.7352
26.2363 | -100.7434
26.24 | -100.7518
26.2437 | -100.7571
26.2474 | -100.7586
26.2511 | -100.7627
26.2548 | -100.7691
26.2585 | -100.776
26.2622 | -100.7822
26.2658 | -100.7866
26.2695 | -100.7918
26.2732 | -100.7986
26.2769 | -100.8035
26.2806 | -100.8066
26.2843 | -100.8099
26.288 | -100.8121
26.2917 | -100.8139
26.2954 | -100.8147
26.299 | -100.8149
26.3027 | -100.8168
26.3064 | -100.8196
26.3101 | -100.8201
26.3138 | -100.8173
26.3175 | -100.8118
26.3212 | -100.8142
26.3249 | -100.8178
26.3286 | -100.8183
26.3322 | -100.8181
26.3359 | -100.8154
26.3396 | -100.8115
26.3433 | -100.8092
26.347 | -100.8087
26.3507 | -100.8066
26.3544 | -100.8016
26.3581 | -100.7948
26.3618 | -100.7879
26.3654 | -100.7771
26.3691 | -100.7703
26.3728 | -100.762
26.3765 | -100.75
26.3802 | -100.7414
26.3839 | -100.7323
26.3876 | -100.7247
26.3913 | -100.7199
26.3949 | -100.7116
26.3986 | -100.6999
26.4023 | -100.6826
26.406 | -100.6664
26.4097 | -100.6528
26.4134 | -100.6396
26.4171 | -100.6243
26.4208 | -100.6094
26.4245 | -100.5959
26.4281 | -100.5821
26.4318 | -100.5677
26.4355 | -100.5538
26.4392 | -100.5415
26.4429 | -100.5291
26.4466 | -100.5167
26.4503 | -100.5033
26.454 | -100.4893
26.4577 | -100.4747
26.4613 | -100.4601
26.465 | -100.4444
26.4687 | -100.4277
26.4724 | -100.4127
26.4761 | -100.3974
26.4798 | -100.3824
26.4835 | -100.3679
26.4872 | -100.3536
26.4908 | -100.3379
26.4945 | -100.3227
26.4982 | -100.3104
26.5019 | -100.2953
26.5056 | -100.2814
26.5093 | -100.269
26.513 | -100.2545
26.5167 | -100.2426
26.5204 | -100.2326
26.524 | -100.2229
26.5277 | -100.213
26.5314 | -100.2032
26.5351 | -100.194
26.5388 | -100.1862
26.5425 | -100.1801
26.5462 | -100.1745
26.5499 | -100.1708
26.5536 | -100.1681
26.5572 | -100.1647
26.5609 | -100.1625
26.5646 | -100.1635
26.5683 | -100.1643
26.572 | -100.1644
26.5757 | -100.1653
26.5794 | -100.1684
26.5831 | -100.1734
26.5867 | -100.1778
26.5904 | -100.1837
26.5941 | -100.1902
26.5978 | -100.1997
26.6015 | -100.2097
26.6052 | -100.2298
26.6089 | -100.2409
26.6126 | -100.2532
26.6163 | -100.2654
26.6199 | -100.2813
26.6236 | -100.2995
26.6273 | -100.3151
26.631 | -100.329
26.6347 | -100.3418
26.6384 | -100.3535
26.6421 | -100.3655
26.6458 | -100.3782
26.6495 | -100.3882
26.6531 | -100.4015
26.6568 | -100.4143
26.6605 | -100.4244
26.6642 | -100.4385
26.6679 | -100.4528
26.6716 | -100.4656
26.6753 | -100.4748
26.679 | -100.4835
26.6826 | -100.4877
26.6863 | -100.4926
26.69 | -100.4967
26.6937 | -100.5039
26.6974 | -100.5074
26.7011 | -100.5051
26.7048 | -100.5033
26.7085 | -100.5013
26.7122 | -100.4967
26.7158 | -100.4912
26.7195 | -100.4829
26.7232 | -100.4733
26.7269 | -100.4634
26.7306 | -100.4528
26.7343 | -100.4405
26.738 | -100.4255
26.7417 | -100.4109
26.7454 | -100.3893
26.749 | -100.3696
26.7527 | -100.3495
26.7564 | -100.3233
26.7601 | -100.2941
26.7638 | -100.263
26.7675 | -100.2319
26.7712 | -100.1988
26.7749 | -100.1617
26.7786 | -100.128
26.7822 | -100.0976
26.7859 | -100.0639
26.7896 | -100.0245
26.7933 | -99.9856
26.797 | -99.9445
26.8007 | -99.8996
26.8044 | -99.8535
26.8081 | -99.8082
26.8117 | -99.7589
26.8154 | -99.7111
26.8191 | -99.665
26.8228 | -99.6189
26.8265 | -99.5692
26.8302 | -99.5189
26.8339 | -99.4744
26.8376 | -99.4293
26.8413 | -99.3839
26.8449 | -99.3393
26.8486 | -99.2957
26.8523 | -99.2509
26.856 | -99.2048
26.8597 | -99.1583
26.8634 | -99.1142
26.8671 | -99.0716
26.8708 | -99.0299
26.8745 | -98.9883
26.8781 | -98.9492
26.8818 | -98.9081
26.8855 | -98.8647
26.8892 | -98.8254
26.8929 | -98.7871
26.8966 | -98.749
26.9003 | -98.711
26.904 | -98.6747
26.9076 | -98.6396
26.9113 | -98.6043
26.915 | -98.5703
26.9187 | -98.5384
26.9224 | -98.5059
26.9261 | -98.4761
26.9298 | -98.4477
26.9335 | -98.4226
26.9372 | -98.3973
26.9408 | -98.3727
26.9445 | -98.3511
26.9482 | -98.3305
26.9519 | -98.3107
26.9556 | -98.2928
26.9593 | -98.2779
26.963 | -98.2634
26.9667 | -98.2512
26.9704 | -98.2416
26.974 | -98.2334
26.9777 | -98.2269
26.9814 | -98.222
26.9851 | -98.2197
26.9888 | -98.2199
26.9925 | -98.2225
26.9962 | -98.2269
26.9999 | -98.2326
27.0036 | -98.2405
27.0072 | -98.2516
27.0109 | -98.2657
27.0146 | -98.2821
27.0183 | -98.3004
27.022 | -98.3217
27.0257 | -98.3442
27.0294 | -98.3686
27.0331 | -98.3924
27.0367 | -98.4187
27.0404 | -98.4477
27.0441 | -98.4758
27.0478 | -98.5047
27.0515 | -98.536
27.0552 | -98.571
27.0589 | -98.6062
27.0626 | -98.6443
27.0663 | -98.6817
27.0699 | -98.7221
27.0736 | -98.7622
27.0773 | -98.801
27.081 | -98.8409
27.0847 | -98.8764
27.0884 | -98.9101
27.0921 | -98.9435
27.0958 | -98.9724
27.0995 | -99.0059
27.1031 | -99.043
27.1068 | -99.078
27.1105 | -99.1134
27.1142 | -99.149
27.1179 | -99.184
27.1216 | -99.2189
27.1253 | -99.2531
27.129 | -99.2847
27.1326 | -99.3152
27.1363 | -99.3262
27.14 | -99.3534
27.1437 | -99.379
27.1474 | -99.4028
27.1511 | -99.4233
27.1548 | -99.4429
27.1585 | -99.463
27.1622 | -99.4801
27.1658 | -99.4896
27.1695 | -99.4972
27.1732 | -99.506
27.1769 | -99.5108
27.1806 | -99.5137
27.1843 | -99.5156
27.188 | -99.5165
27.1917 | -99.5152
27.1954 | -99.5154
27.199 | -99.5092
27.2027 | -99.4996
27.2064 | -99.4903
27.2101 | -99.4834
27.2138 | -99.4755
27.2175 | -99.4601
27.2212 | -99.4455
27.2249 | -99.4327
27.2285 | -99.4187
27.2322 | -99.4042
27.2359 | -99.391
27.2396 | -99.3796
27.2433 | -99.3695
27.247 | -99.3575
27.2507 | -99.3445
27.2544 | -99.3318
27.2581 | -99.3223
27.2617 | -99.3145
27.2654 | -99.3085
27.2691 | -99.3043
27.2728 | -99.2983
27.2765 | -99.2926
27.2802 | -99.2895
27.2839 | -99.2895
27.2876 | -99.2911
27.2913 | -99.2941
27.2949 | -99.3004
27.2986 | -99.3076
27.3023 | -99.3157
27.306 | -99.3273
27.3097 | -99.3383
27.3134 | -99.3499
27.3171 | -99.3615
27.3208 | -99.376
27.3244 | -99.3912
27.3281 | -99.4075
27.3318 | -99.4261
27.3355 | -99.4473
27.3392 | -99.4697
27.3429 | -99.4911
27.3466 | -99.5131
27.3503 | -99.535
27.354 | -99.5573
27.3576 | -99.58
27.3613 | -99.6041
27.365 | -99.6295
27.3687 | -99.6541
27.3724 | -99.6781
27.3761 | -99.7025
27.3798 | -99.7281
27.3835 | -99.753
27.3872 | -99.7784
27.3908 | -99.802
27.3945 | -99.8231
27.3982 | -99.8441
27.4019 | -99.8667
27.4056 | -99.8881
27.4093 | -99.9093
27.413 | -99.9309
27.4167 | -99.9524
27.4204 | -99.9726
27.424 | -99.9927
27.4277 | -100.0129
27.4314 | -100.0322
27.4351 | -100.0499
27.4388 | -100.0648
27.4425 | -100.081
27.4462 | -100.097
27.4499 | -100.1129
27.4535 | -100.1268
27.4572 | -100.1408
27.4609 | -100.1551
27.4646 | -100.1689
27.4683 | -100.1869
27.472 | -100.2052
27.4757 | -100.2223
27.4794 | -100.2414
27.4831 | -100.2541
27.4867 | -100.2612
27.4904 | -100.2691
27.4941 | -100.2755
27.4978 | -100.2821
27.5015 | -100.2886
27.5052 | -100.2937
27.5089 | -100.2991
27.5126 | -100.3038
27.5163 | -100.3073
27.5199 | -100.3101
27.5236 | -100.3137
27.5273 | -100.3195
27.531 | -100.3261
27.5347 | -100.3309
27.5384 | -100.3324
27.5421 | -100.3303
27.5458 | -100.3262
27.5494 | -100.3203
27.5531 | -100.3139
27.5568 | -100.3074
27.5605 | -100.3063
27.5642 | -100.3069
27.5679 | -100.3071
27.5716 | -100.3064
27.5753 | -100.3048
27.579 | -100.3018
27.5826 | -100.299
27.5863 | -100.2976
27.59 | -100.2956
27.5937 | -100.2918
27.5974 | -100.2868
27.6011 | -100.2802
27.6048 | -100.2733
27.6085 | -100.2665
27.6122 | -100.258
27.6158 | -100.2487
27.6195 | -100.2383
27.6232 | -100.229
27.6269 | -100.2188
27.6306 | -100.2064
27.6343 | -100.1928
27.638 | -100.1813
27.6417 | -100.1684
27.6454 | -100.152
27.649 | -100.1363
27.6527 | -100.12
27.6564 | -100.1048
27.6601 | -100.0911
27.6638 | -100.0763
27.6675 | -100.0619
27.6712 | -100.046
27.6749 | -100.0289
27.6785 | -100.013
27.6822 | -99.9967
27.6859 | -99.9807
27.6896 | -99.9672
27.6933 | -99.9535
27.697 | -99.9396
27.7007 | -99.9251
27.7044 | -99.913
27.7081 | -99.9026
27.7117 | -99.8942
27.7154 | -99.8878
27.7191 | -99.8806
27.7228 | -99.8752
27.7265 | -99.8702
27.7302 | -99.8658
27.7339 | -99.8638
27.7376 | -99.8649
27.7413 | -99.8669
27.7449 | -99.8687
27.7486 | -99.8731
27.7523 | -99.8787
27.756 | -99.8854
27.7597 | -99.8918
27.7634 | -99.899
27.7671 | -99.9074
27.7708 | -99.9169
27.7744 | -99.9266
27.7781 | -99.9381
27.7818 | -99.9503
27.7855 | -99.9658
27.7892 | -99.9826
27.7929 | -99.9998
27.7966 | -100.0177
27.8003 | -100.0342
27.804 | -100.0509
27.8076 | -100.0683
27.8113 | -100.0873
27.815 | -100.1071
27.8187 | -100.125
27.8224 | -100.1445
27.8261 | -100.1641
27.8298 | -100.1855
27.8335 | -100.204
27.8372 | -100.2219
27.8408 | -100.242
27.8445 | -100.2619
27.8482 | -100.2802
27.8519 | -100.2996
27.8556 | -100.319
27.8593 | -100.3395
27.863 | -100.3601
27.8667 | -100.3785
27.8703 | -100.3976
27.874 | -100.4169
27.8777 | -100.4353
27.8814 | -100.4539
27.8851 | -100.4718
27.8888 | -100.4893
27.8925 | -100.508
27.8962 | -100.5274
27.8999 | -100.5457
27.9035 | -100.5651
27.9072 | -100.5846
27.9109 | -100.6016
27.9146 | -100.6173
27.9183 | -100.6329
27.922 | -100.6478
27.9257 | -100.6591
27.9294 | -100.6734
27.9331 | -100.6929
27.9367 | -100.713
27.9404 | -100.7297
27.9441 | -100.7448
27.9478 | -100.7561
27.9515 | -100.76
27.9552 | -100.7613
27.9589 | -100.7622
27.9626 | -100.762
27.9662 | -100.7627
27.9699 | -100.7635
27.9736 | -100.7625
27.9773 | -100.7618
27.981 | -100.7593
27.9847 | -100.7546
27.9884 | -100.7518
27.9921 | -100.7479
27.9958 | -100.7407
27.9994 | -100.7282
28.0031 | -100.716