## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 249 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 267
Hood | 02-92B/P | 193
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1010.8 | hPa |  | PRC PRO 01
Humidity | 53.8 | %/Hr
Internal temperature of calibration bench | 23.88 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1018 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.5 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.99 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3103 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 15/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 21/06/2018
Technical Manager | Anthony HUE | Approval |  | 28/06/2018
CEO | Robert GLEMOT | Approval |  | 28/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4366 | 57.3944
0.012209 | -36.0102 | 38.2616
0.018313 | -34.9533 | 28.1484
0.024417 | -34.492 | 21.8994
0.030521 | -34.3821 | 18.1451
0.036626 | -34.2394 | 15.1013
0.04273 | -34.1448 | 12.7449
0.048834 | -34.1001 | 11.2328
0.054938 | -34.0357 | 10.097
0.061043 | -33.9751 | 9.0117
0.067147 | -33.9644 | 8.1903
0.073251 | -33.9651 | 7.4314
0.079355 | -33.9834 | 6.7744
0.08546 | -33.9978 | 6.1862
0.091564 | -33.9932 | 5.7461
0.097668 | -33.9749 | 5.4928
0.10377 | -33.9524 | 5.246
0.10988 | -33.9345 | 4.9704
0.11598 | -33.9243 | 4.6807
0.12209 | -33.9252 | 4.4003
0.12819 | -33.935 | 4.1195
0.13429 | -33.9466 | 3.8329
0.1404 | -33.951 | 3.5517
0.1465 | -33.9536 | 3.2907
0.15261 | -33.9512 | 3.0666
0.15871 | -33.9447 | 2.9013
0.16482 | -33.9361 | 2.7974
0.17092 | -33.9333 | 2.761
0.17702 | -33.9308 | 2.6908
0.18313 | -33.9296 | 2.6152
0.18923 | -33.929 | 2.5091
0.19534 | -33.9276 | 2.3848
0.20144 | -33.9264 | 2.2444
0.20754 | -33.9236 | 2.1248
0.21365 | -33.9226 | 2.0115
0.21975 | -33.9222 | 1.9148
0.22586 | -33.9233 | 1.8285
0.23196 | -33.9257 | 1.7568
0.23807 | -33.9298 | 1.6946
0.24417 | -33.9339 | 1.6366
0.25027 | -33.938 | 1.5749
0.25638 | -33.9405 | 1.5139
0.26248 | -33.9425 | 1.4481
0.26859 | -33.9436 | 1.384
0.27469 | -33.9442 | 1.3142
0.2808 | -33.9454 | 1.2419
0.2869 | -33.9464 | 1.1722
0.293 | -33.9474 | 1.1061
0.29911 | -33.9484 | 1.0469
0.30521 | -33.9494 | 0.98834
0.31132 | -33.9499 | 0.94992
0.31742 | -33.9511 | 0.90756
0.32353 | -33.9512 | 0.87292
0.32963 | -33.9519 | 0.84423
0.33573 | -33.9527 | 0.82431
0.34184 | -33.9539 | 0.80343
0.34794 | -33.9558 | 0.7802
0.35405 | -33.9578 | 0.75661
0.36015 | -33.9601 | 0.73167
0.36626 | -33.962 | 0.70433
0.37236 | -33.964 | 0.66775
0.37846 | -33.9662 | 0.62359
0.38457 | -33.9671 | 0.56607
0.39067 | -33.9682 | 0.51926
0.39678 | -33.9689 | 0.46714
0.40288 | -33.9694 | 0.41281
0.40899 | -33.9694 | 0.36055
0.41509 | -33.9687 | 0.30944
0.42119 | -33.9687 | 0.26617
0.4273 | -33.9687 | 0.21988
0.4334 | -33.9686 | 0.17381
0.43951 | -33.9689 | 0.1306
0.44561 | -33.9697 | 0.09109
0.45172 | -33.9712 | 0.059609
0.45782 | -33.972 | 0.022729
0.46392 | -33.9732 | -0.013325
0.47003 | -33.9742 | -0.045473
0.47613 | -33.9751 | -0.079416
0.48224 | -33.9756 | -0.11408
0.48834 | -33.9747 | -0.14895
0.49445 | -33.9745 | -0.18422
0.50055 | -33.9739 | -0.21636
0.50665 | -33.9732 | -0.2507
0.51276 | -33.9722 | -0.28494
0.51886 | -33.9712 | -0.319
0.52497 | -33.97 | -0.3434
0.53107 | -33.9693 | -0.37308
0.53717 | -33.9687 | -0.40133
0.54328 | -33.9678 | -0.42429
0.54938 | -33.9677 | -0.44691
0.55549 | -33.9677 | -0.46797
0.56159 | -33.9688 | -0.4797
0.5677 | -33.9692 | -0.49443
0.5738 | -33.9702 | -0.51083
0.5799 | -33.971 | -0.52587
0.58601 | -33.9721 | -0.53608
0.59211 | -33.9737 | -0.54244
0.59822 | -33.9751 | -0.55199
0.60432 | -33.9766 | -0.56413
0.61043 | -33.9778 | -0.57213
0.61653 | -33.979 | -0.58246
0.62263 | -33.9805 | -0.59184
0.62874 | -33.9811 | -0.60241
0.63484 | -33.9819 | -0.61168
0.64095 | -33.9824 | -0.62193
0.64705 | -33.9828 | -0.63162
0.65316 | -33.983 | -0.64395
0.65926 | -33.983 | -0.65968
0.66536 | -33.9819 | -0.67694
0.67147 | -33.9815 | -0.69009
0.67757 | -33.9811 | -0.7043
0.68368 | -33.9804 | -0.72316
0.68978 | -33.9796 | -0.74026
0.69589 | -33.9787 | -0.75922
0.70199 | -33.9777 | -0.78191
0.70809 | -33.9768 | -0.80108
0.7142 | -33.9761 | -0.82146
0.7203 | -33.9751 | -0.84068
0.72641 | -33.9744 | -0.86313
0.73251 | -33.9738 | -0.88313
0.73862 | -33.974 | -0.90494
0.74472 | -33.9737 | -0.92489
0.75082 | -33.9734 | -0.94704
0.75693 | -33.9737 | -0.96881
0.76303 | -33.9743 | -0.99072
0.76914 | -33.9761 | -1.0111
0.77524 | -33.9771 | -1.0307
0.78135 | -33.9784 | -1.0507
0.78745 | -33.9797 | -1.0705
0.79355 | -33.981 | -1.0902
0.79966 | -33.9827 | -1.1095
0.80576 | -33.985 | -1.129
0.81187 | -33.9866 | -1.1505
0.81797 | -33.9885 | -1.1716
0.82408 | -33.9902 | -1.192
0.83018 | -33.9919 | -1.2125
0.83628 | -33.9938 | -1.2308
0.84239 | -33.9949 | -1.2514
0.84849 | -33.9963 | -1.273
0.8546 | -33.9978 | -1.2928
0.8607 | -33.999 | -1.3135
0.86681 | -34.0001 | -1.3324
0.87291 | -34.001 | -1.3512
0.87901 | -34.0015 | -1.3693
0.88512 | -34.0022 | -1.3874
0.89122 | -34.0029 | -1.4057
0.89733 | -34.0033 | -1.4214
0.90343 | -34.0037 | -1.4391
0.90953 | -34.0039 | -1.4535
0.91564 | -34.0039 | -1.4694
0.92174 | -34.0041 | -1.4845
0.92785 | -34.0042 | -1.4991
0.93395 | -34.0043 | -1.513
0.94006 | -34.0045 | -1.526
0.94616 | -34.0042 | -1.5379
0.95226 | -34.0043 | -1.5506
0.95837 | -34.0044 | -1.5646
0.96447 | -34.0046 | -1.5775
0.97058 | -34.0045 | -1.5909
0.97668 | -34.0043 | -1.6044
0.98279 | -34.004 | -1.6172
0.98889 | -34.0039 | -1.6279
0.99499 | -34.0038 | -1.6413
1.0011 | -34.0037 | -1.6521
1.0072 | -34.0034 | -1.6637
1.0133 | -34.0031 | -1.6745
1.0194 | -34.0027 | -1.6875
1.0255 | -34.0024 | -1.7012
1.0316 | -34.0018 | -1.7132
1.0377 | -34.0014 | -1.7269
1.0438 | -34.0009 | -1.7396
1.0499 | -34.0001 | -1.7549
1.056 | -33.9994 | -1.7692
1.0621 | -33.9988 | -1.7835
1.0682 | -33.9984 | -1.7959
1.0743 | -33.9977 | -1.8114
1.0805 | -33.9972 | -1.8258
1.0866 | -33.9963 | -1.8409
1.0927 | -33.9955 | -1.8554
1.0988 | -33.9949 | -1.8703
1.1049 | -33.994 | -1.8845
1.111 | -33.9933 | -1.8998
1.1171 | -33.9925 | -1.9144
1.1232 | -33.9921 | -1.9296
1.1293 | -33.9917 | -1.9436
1.1354 | -33.9914 | -1.9586
1.1415 | -33.991 | -1.9742
1.1476 | -33.9904 | -1.9898
1.1537 | -33.9898 | -2.0045
1.1598 | -33.9899 | -2.0193
1.1659 | -33.9897 | -2.0326
1.172 | -33.9894 | -2.0466
1.1781 | -33.9892 | -2.0605
1.1842 | -33.9891 | -2.0748
1.1903 | -33.9894 | -2.0858
1.1964 | -33.9893 | -2.0992
1.2025 | -33.9894 | -2.1132
1.2086 | -33.9897 | -2.1257
1.2147 | -33.9898 | -2.1381
1.2209 | -33.9901 | -2.1494
1.227 | -33.9908 | -2.1602
1.2331 | -33.9912 | -2.1722
1.2392 | -33.9916 | -2.1841
1.2453 | -33.992 | -2.1952
1.2514 | -33.9924 | -2.2052
1.2575 | -33.993 | -2.2162
1.2636 | -33.9939 | -2.2252
1.2697 | -33.9945 | -2.2344
1.2758 | -33.9951 | -2.2456
1.2819 | -33.9959 | -2.2556
1.288 | -33.9965 | -2.2664
1.2941 | -33.9974 | -2.2749
1.3002 | -33.998 | -2.2829
1.3063 | -33.9986 | -2.292
1.3124 | -33.9992 | -2.3011
1.3185 | -34 | -2.3102
1.3246 | -34.0007 | -2.3193
1.3307 | -34.0015 | -2.3279
1.3368 | -34.0023 | -2.3366
1.3429 | -34.0029 | -2.3455
1.349 | -34.0034 | -2.3549
1.3551 | -34.0041 | -2.3642
1.3613 | -34.0047 | -2.3725
1.3674 | -34.0053 | -2.3825
1.3735 | -34.0057 | -2.3922
1.3796 | -34.0062 | -2.4013
1.3857 | -34.0065 | -2.4105
1.3918 | -34.0071 | -2.4206
1.3979 | -34.0076 | -2.4297
1.404 | -34.0081 | -2.4409
1.4101 | -34.0085 | -2.4512
1.4162 | -34.0089 | -2.4617
1.4223 | -34.0094 | -2.473
1.4284 | -34.0099 | -2.4829
1.4345 | -34.0101 | -2.4943
1.4406 | -34.0104 | -2.5077
1.4467 | -34.0109 | -2.5192
1.4528 | -34.0113 | -2.5304
1.4589 | -34.0118 | -2.5422
1.465 | -34.0121 | -2.5529
1.4711 | -34.0125 | -2.5641
1.4772 | -34.0131 | -2.5773
1.4833 | -34.0135 | -2.5901
1.4894 | -34.014 | -2.6022
1.4955 | -34.0146 | -2.614
1.5016 | -34.0151 | -2.6264
1.5078 | -34.0155 | -2.6401
1.5139 | -34.0159 | -2.653
1.52 | -34.0163 | -2.6654
1.5261 | -34.0167 | -2.6779
1.5322 | -34.017 | -2.6917
1.5383 | -34.0174 | -2.7041
1.5444 | -34.0179 | -2.7166
1.5505 | -34.0184 | -2.7298
1.5566 | -34.0189 | -2.7427
1.5627 | -34.0193 | -2.7563
1.5688 | -34.0199 | -2.7697
1.5749 | -34.0203 | -2.7821
1.581 | -34.0206 | -2.7953
1.5871 | -34.021 | -2.8089
1.5932 | -34.0213 | -2.8216
1.5993 | -34.0217 | -2.8357
1.6054 | -34.0221 | -2.8488
1.6115 | -34.0225 | -2.8627
1.6176 | -34.0228 | -2.8752
1.6237 | -34.023 | -2.8886
1.6298 | -34.0232 | -2.901
1.6359 | -34.0236 | -2.9136
1.642 | -34.024 | -2.9254
1.6482 | -34.0242 | -2.9383
1.6543 | -34.0245 | -2.9518
1.6604 | -34.0249 | -2.9637
1.6665 | -34.0249 | -2.976
1.6726 | -34.0251 | -2.9884
1.6787 | -34.0255 | -3.0008
1.6848 | -34.0254 | -3.0133
1.6909 | -34.0256 | -3.0259
1.697 | -34.0258 | -3.0379
1.7031 | -34.0259 | -3.0504
1.7092 | -34.0258 | -3.0624
1.7153 | -34.026 | -3.0743
1.7214 | -34.0258 | -3.0864
1.7275 | -34.0257 | -3.0989
1.7336 | -34.0256 | -3.1118
1.7397 | -34.0258 | -3.124
1.7458 | -34.0257 | -3.1361
1.7519 | -34.0257 | -3.1482
1.758 | -34.0258 | -3.1601
1.7641 | -34.0259 | -3.1716
1.7702 | -34.0258 | -3.1845
1.7763 | -34.0258 | -3.1959
1.7824 | -34.0257 | -3.2082
1.7885 | -34.0257 | -3.2206
1.7947 | -34.0258 | -3.2326
1.8008 | -34.0258 | -3.2438
1.8069 | -34.0258 | -3.2552
1.813 | -34.0258 | -3.2674
1.8191 | -34.0258 | -3.28
1.8252 | -34.0258 | -3.2927
1.8313 | -34.026 | -3.3049
1.8374 | -34.0261 | -3.3167
1.8435 | -34.0261 | -3.3285
1.8496 | -34.0262 | -3.3416
1.8557 | -34.0263 | -3.3543
1.8618 | -34.0266 | -3.3665
1.8679 | -34.0267 | -3.3784
1.874 | -34.0267 | -3.3904
1.8801 | -34.0269 | -3.4028
1.8862 | -34.0271 | -3.4146
1.8923 | -34.0272 | -3.4268
1.8984 | -34.0273 | -3.4388
1.9045 | -34.0274 | -3.4504
1.9106 | -34.0274 | -3.4626
1.9167 | -34.0275 | -3.4747
1.9228 | -34.0276 | -3.4865
1.9289 | -34.0277 | -3.4984
1.9351 | -34.0278 | -3.5092
1.9412 | -34.0281 | -3.5204
1.9473 | -34.0283 | -3.5324
1.9534 | -34.0285 | -3.5439
1.9595 | -34.0287 | -3.5556
1.9656 | -34.0289 | -3.568
1.9717 | -34.0292 | -3.5796
1.9778 | -34.0295 | -3.591
1.9839 | -34.0297 | -3.6025
1.99 | -34.0299 | -3.6136
1.9961 | -34.03 | -3.6257
2.0022 | -34.0302 | -3.6371
2.0083 | -34.0303 | -3.6483
2.0144 | -34.0305 | -3.6597
2.0205 | -34.0308 | -3.6714
2.0266 | -34.0311 | -3.6829
2.0327 | -34.0313 | -3.6945
2.0388 | -34.0318 | -3.706
2.0449 | -34.0319 | -3.7174
2.051 | -34.0322 | -3.7288
2.0571 | -34.0325 | -3.7401
2.0632 | -34.0327 | -3.7516
2.0693 | -34.0329 | -3.7638
2.0754 | -34.0332 | -3.7751
2.0816 | -34.0335 | -3.7873
2.0877 | -34.0337 | -3.7989
2.0938 | -34.034 | -3.811
2.0999 | -34.0342 | -3.8227
2.106 | -34.0344 | -3.8343
2.1121 | -34.0346 | -3.8463
2.1182 | -34.0348 | -3.8571
2.1243 | -34.0351 | -3.8679
2.1304 | -34.0353 | -3.8791
2.1365 | -34.0356 | -3.8908
2.1426 | -34.0358 | -3.9028
2.1487 | -34.0361 | -3.9141
2.1548 | -34.0365 | -3.926
2.1609 | -34.037 | -3.9374
2.167 | -34.0373 | -3.9486
2.1731 | -34.0376 | -3.9607
2.1792 | -34.0378 | -3.9718
2.1853 | -34.0382 | -3.9838
2.1914 | -34.0385 | -3.9952
2.1975 | -34.0389 | -4.0068
2.2036 | -34.0391 | -4.0189
2.2097 | -34.0394 | -4.0304
2.2158 | -34.0397 | -4.0411
2.222 | -34.04 | -4.0516
2.2281 | -34.0402 | -4.0631
2.2342 | -34.0406 | -4.0745
2.2403 | -34.0407 | -4.0852
2.2464 | -34.041 | -4.0959
2.2525 | -34.0413 | -4.1074
2.2586 | -34.0416 | -4.118
2.2647 | -34.0419 | -4.1285
2.2708 | -34.0422 | -4.1397
2.2769 | -34.0425 | -4.1505
2.283 | -34.0427 | -4.1617
2.2891 | -34.043 | -4.1726
2.2952 | -34.0431 | -4.1832
2.3013 | -34.0434 | -4.1941
2.3074 | -34.0436 | -4.2046
2.3135 | -34.0438 | -4.216
2.3196 | -34.044 | -4.2262
2.3257 | -34.0442 | -4.237
2.3318 | -34.0444 | -4.2481
2.3379 | -34.0446 | -4.2589
2.344 | -34.0447 | -4.2704
2.3501 | -34.0449 | -4.2811
2.3562 | -34.0451 | -4.2915
2.3623 | -34.0452 | -4.3019
2.3685 | -34.0453 | -4.3125
2.3746 | -34.0454 | -4.3237
2.3807 | -34.0457 | -4.3342
2.3868 | -34.0458 | -4.3449
2.3929 | -34.046 | -4.3556
2.399 | -34.0461 | -4.3661
2.4051 | -34.0464 | -4.3767
2.4112 | -34.0465 | -4.3873
2.4173 | -34.0467 | -4.398
2.4234 | -34.0468 | -4.4088
2.4295 | -34.0469 | -4.4196
2.4356 | -34.0471 | -4.4305
2.4417 | -34.0472 | -4.4416
2.4478 | -34.0473 | -4.4525
2.4539 | -34.0475 | -4.4627
2.46 | -34.0476 | -4.4741
2.4661 | -34.0478 | -4.4852
2.4722 | -34.0479 | -4.4954
2.4783 | -34.048 | -4.5058
2.4844 | -34.0481 | -4.5169
2.4905 | -34.0482 | -4.5278
2.4966 | -34.0483 | -4.5386
2.5027 | -34.0484 | -4.549
2.5089 | -34.0485 | -4.5599
2.515 | -34.0486 | -4.5707
2.5211 | -34.0487 | -4.5814
2.5272 | -34.0487 | -4.5923
2.5333 | -34.0488 | -4.6032
2.5394 | -34.049 | -4.614
2.5455 | -34.0491 | -4.6253
2.5516 | -34.0493 | -4.6356
2.5577 | -34.0494 | -4.6471
2.5638 | -34.0496 | -4.6587
2.5699 | -34.0496 | -4.6701
2.576 | -34.0497 | -4.6816
2.5821 | -34.0498 | -4.6929
2.5882 | -34.0499 | -4.7039
2.5943 | -34.05 | -4.7147
2.6004 | -34.0502 | -4.7257
2.6065 | -34.0503 | -4.736
2.6126 | -34.0504 | -4.7464
2.6187 | -34.0505 | -4.7568
2.6248 | -34.0506 | -4.7676
2.6309 | -34.0508 | -4.779
2.637 | -34.051 | -4.7905
2.6431 | -34.0512 | -4.801
2.6492 | -34.0513 | -4.8123
2.6554 | -34.0515 | -4.8231
2.6615 | -34.0515 | -4.834
2.6676 | -34.0518 | -4.8454
2.6737 | -34.052 | -4.8565
2.6798 | -34.0523 | -4.8675
2.6859 | -34.0525 | -4.8788
2.692 | -34.0526 | -4.8896
2.6981 | -34.0528 | -4.9007
2.7042 | -34.053 | -4.9118
2.7103 | -34.0532 | -4.9224
2.7164 | -34.0535 | -4.9333
2.7225 | -34.0537 | -4.9446
2.7286 | -34.0539 | -4.9556
2.7347 | -34.0541 | -4.9661
2.7408 | -34.0544 | -4.9761
2.7469 | -34.0546 | -4.9874
2.753 | -34.0548 | -4.9979
2.7591 | -34.0551 | -5.0092
2.7652 | -34.0554 | -5.0206
2.7713 | -34.0555 | -5.0313
2.7774 | -34.0558 | -5.0422
2.7835 | -34.0559 | -5.0531
2.7896 | -34.0561 | -5.0636
2.7958 | -34.0564 | -5.0743
2.8019 | -34.0566 | -5.0847
2.808 | -34.0569 | -5.0954
2.8141 | -34.0572 | -5.1065
2.8202 | -34.0575 | -5.1173
2.8263 | -34.0578 | -5.1282
2.8324 | -34.0581 | -5.1395
2.8385 | -34.0585 | -5.1503
2.8446 | -34.0588 | -5.161
2.8507 | -34.0592 | -5.1719
2.8568 | -34.0595 | -5.1823
2.8629 | -34.0598 | -5.1929
2.869 | -34.0601 | -5.204
2.8751 | -34.0604 | -5.2147
2.8812 | -34.0608 | -5.2259
2.8873 | -34.0611 | -5.2368
2.8934 | -34.0614 | -5.2478
2.8995 | -34.0617 | -5.2589
2.9056 | -34.062 | -5.27
2.9117 | -34.0623 | -5.281
2.9178 | -34.0627 | -5.2917
2.9239 | -34.0631 | -5.3032
2.93 | -34.0635 | -5.3147
2.9361 | -34.0638 | -5.3254
2.9423 | -34.0641 | -5.3361
2.9484 | -34.0645 | -5.3471
2.9545 | -34.0649 | -5.3581
2.9606 | -34.0653 | -5.3695
2.9667 | -34.0656 | -5.3806
2.9728 | -34.066 | -5.3919
2.9789 | -34.0664 | -5.403
2.985 | -34.0668 | -5.4143
2.9911 | -34.0672 | -5.4256
2.9972 | -34.0677 | -5.4366
3.0033 | -34.068 | -5.4481
3.0094 | -34.0684 | -5.4592
3.0155 | -34.0688 | -5.4701
3.0216 | -34.0692 | -5.4813
3.0277 | -34.0695 | -5.493
3.0338 | -34.0699 | -5.5044
3.0399 | -34.0703 | -5.5157
3.046 | -34.0706 | -5.5272
3.0521 | -34.071 | -5.5384
3.0582 | -34.0713 | -5.5504
3.0643 | -34.0716 | -5.5619
3.0704 | -34.0719 | -5.5737
3.0765 | -34.0722 | -5.5848
3.0827 | -34.0725 | -5.5961
3.0888 | -34.0729 | -5.6073
3.0949 | -34.0733 | -5.6187
3.101 | -34.0736 | -5.6302
3.1071 | -34.0739 | -5.6419
3.1132 | -34.0743 | -5.6537
3.1193 | -34.0746 | -5.6652
3.1254 | -34.075 | -5.6765
3.1315 | -34.0754 | -5.6884
3.1376 | -34.0757 | -5.6998
3.1437 | -34.076 | -5.7114
3.1498 | -34.0764 | -5.7235
3.1559 | -34.0767 | -5.7349
3.162 | -34.0769 | -5.747
3.1681 | -34.0772 | -5.759
3.1742 | -34.0775 | -5.7706
3.1803 | -34.0779 | -5.7822
3.1864 | -34.0782 | -5.7935
3.1925 | -34.0785 | -5.8051
3.1986 | -34.0788 | -5.8171
3.2047 | -34.0791 | -5.8292
3.2108 | -34.0793 | -5.8403
3.2169 | -34.0796 | -5.852
3.223 | -34.0799 | -5.8638
3.2292 | -34.0801 | -5.8756
3.2353 | -34.0803 | -5.8865
3.2414 | -34.0806 | -5.8979
3.2475 | -34.0808 | -5.9093
3.2536 | -34.0811 | -5.9208
3.2597 | -34.0813 | -5.9318
3.2658 | -34.0815 | -5.9427
3.2719 | -34.0818 | -5.9538
3.278 | -34.082 | -5.9654
3.2841 | -34.0823 | -5.9765
3.2902 | -34.0825 | -5.9879
3.2963 | -34.0828 | -5.9994
3.3024 | -34.0829 | -6.0105
3.3085 | -34.0832 | -6.0221
3.3146 | -34.0833 | -6.0333
3.3207 | -34.0836 | -6.0447
3.3268 | -34.0838 | -6.0562
3.3329 | -34.084 | -6.067
3.339 | -34.0841 | -6.0778
3.3451 | -34.0843 | -6.0891
3.3512 | -34.0845 | -6.0998
3.3573 | -34.0847 | -6.111
3.3634 | -34.0849 | -6.1222
3.3696 | -34.0851 | -6.1336
3.3757 | -34.0852 | -6.1444
3.3818 | -34.0855 | -6.1553
3.3879 | -34.0856 | -6.1664
3.394 | -34.0858 | -6.1777
3.4001 | -34.086 | -6.1886
3.4062 | -34.0861 | -6.1995
3.4123 | -34.0862 | -6.2101
3.4184 | -34.0864 | -6.2214
3.4245 | -34.0866 | -6.2322
3.4306 | -34.0867 | -6.2429
3.4367 | -34.0869 | -6.2535
3.4428 | -34.087 | -6.2639
3.4489 | -34.0872 | -6.2749
3.455 | -34.0873 | -6.2857
3.4611 | -34.0875 | -6.297
3.4672 | -34.0876 | -6.308
3.4733 | -34.0878 | -6.3191
3.4794 | -34.0879 | -6.3299
3.4855 | -34.0881 | -6.3403
3.4916 | -34.0883 | -6.3507
3.4977 | -34.0884 | -6.362
3.5038 | -34.0886 | -6.3729
3.5099 | -34.0887 | -6.3832
3.5161 | -34.0888 | -6.3935
3.5222 | -34.0889 | -6.4039
3.5283 | -34.089 | -6.4147
3.5344 | -34.0892 | -6.4255
3.5405 | -34.0893 | -6.4361
3.5466 | -34.0895 | -6.447
3.5527 | -34.0896 | -6.4578
3.5588 | -34.0897 | -6.4684
3.5649 | -34.0897 | -6.4791
3.571 | -34.0899 | -6.4897
3.5771 | -34.09 | -6.5008
3.5832 | -34.0901 | -6.5107
3.5893 | -34.0903 | -6.5214
3.5954 | -34.0905 | -6.5322
3.6015 | -34.0907 | -6.5431
3.6076 | -34.0909 | -6.5536
3.6137 | -34.091 | -6.5637
3.6198 | -34.0913 | -6.574
3.6259 | -34.0914 | -6.5849
3.632 | -34.0915 | -6.5955
3.6381 | -34.0917 | -6.6063
3.6442 | -34.0918 | -6.6168
3.6503 | -34.092 | -6.627
3.6565 | -34.0923 | -6.6378
3.6626 | -34.0924 | -6.6484
3.6687 | -34.0926 | -6.6591
3.6748 | -34.0928 | -6.6695
3.6809 | -34.0931 | -6.6802
3.687 | -34.0932 | -6.6911
3.6931 | -34.0935 | -6.7014
3.6992 | -34.0937 | -6.712
3.7053 | -34.0939 | -6.7225
3.7114 | -34.094 | -6.7332
3.7175 | -34.0942 | -6.7439
3.7236 | -34.0945 | -6.7539
3.7297 | -34.0947 | -6.7647
3.7358 | -34.0949 | -6.7756
3.7419 | -34.0951 | -6.7865
3.748 | -34.0953 | -6.797
3.7541 | -34.0955 | -6.8076
3.7602 | -34.0958 | -6.8181
3.7663 | -34.0959 | -6.8286
3.7724 | -34.0962 | -6.8392
3.7785 | -34.0965 | -6.8496
3.7846 | -34.0966 | -6.8601
3.7907 | -34.0969 | -6.871
3.7969 | -34.0972 | -6.8818
3.803 | -34.0974 | -6.892
3.8091 | -34.0977 | -6.9025
3.8152 | -34.0979 | -6.913
3.8213 | -34.0981 | -6.9239
3.8274 | -34.0984 | -6.9342
3.8335 | -34.0987 | -6.9449
3.8396 | -34.099 | -6.9555
3.8457 | -34.0993 | -6.9663
3.8518 | -34.0997 | -6.9768
3.8579 | -34.1 | -6.9873
3.864 | -34.1003 | -6.9981
3.8701 | -34.1006 | -7.0092
3.8762 | -34.1008 | -7.0196
3.8823 | -34.1011 | -7.0304
3.8884 | -34.1014 | -7.0411
3.8945 | -34.1017 | -7.052
3.9006 | -34.102 | -7.063
3.9067 | -34.1023 | -7.074
3.9128 | -34.1026 | -7.0848
3.9189 | -34.1029 | -7.0952
3.925 | -34.1031 | -7.1061
3.9311 | -34.1034 | -7.1163
3.9372 | -34.1037 | -7.1272
3.9434 | -34.104 | -7.1379
3.9495 | -34.1044 | -7.1487
3.9556 | -34.1047 | -7.1592
3.9617 | -34.105 | -7.1695
3.9678 | -34.1053 | -7.1799
3.9739 | -34.1056 | -7.1906
3.98 | -34.1058 | -7.2014
3.9861 | -34.1061 | -7.212
3.9922 | -34.1064 | -7.2227
3.9983 | -34.1067 | -7.2329
4.0044 | -34.107 | -7.2432
4.0105 | -34.1073 | -7.2538
4.0166 | -34.1076 | -7.2646
4.0227 | -34.1079 | -7.2752
4.0288 | -34.1081 | -7.2863
4.0349 | -34.1084 | -7.2967
4.041 | -34.1086 | -7.3073
4.0471 | -34.1089 | -7.3181
4.0532 | -34.1092 | -7.3284
4.0593 | -34.1095 | -7.3392
4.0654 | -34.1098 | -7.3496
4.0715 | -34.1101 | -7.3607
4.0776 | -34.1104 | -7.3713
4.0838 | -34.1107 | -7.3823
4.0899 | -34.111 | -7.3927
4.096 | -34.1112 | -7.4033
4.1021 | -34.1115 | -7.4137
4.1082 | -34.1118 | -7.4241
4.1143 | -34.112 | -7.4344
4.1204 | -34.1123 | -7.4447
4.1265 | -34.1126 | -7.455
4.1326 | -34.1129 | -7.4653
4.1387 | -34.1132 | -7.4757
4.1448 | -34.1135 | -7.4864
4.1509 | -34.1137 | -7.497
4.157 | -34.114 | -7.5079
4.1631 | -34.1142 | -7.5187
4.1692 | -34.1145 | -7.5291
4.1753 | -34.1148 | -7.5396
4.1814 | -34.115 | -7.5498
4.1875 | -34.1153 | -7.5606
4.1936 | -34.1156 | -7.5713
4.1997 | -34.1158 | -7.5815
4.2058 | -34.116 | -7.5923
4.2119 | -34.1163 | -7.6027
4.218 | -34.1165 | -7.613
4.2241 | -34.1168 | -7.6233
4.2303 | -34.1171 | -7.6336
4.2364 | -34.1173 | -7.6441
4.2425 | -34.1175 | -7.6546
4.2486 | -34.1177 | -7.6653
4.2547 | -34.1179 | -7.6755
4.2608 | -34.1181 | -7.6861
4.2669 | -34.1183 | -7.6962
4.273 | -34.1185 | -7.7068
4.2791 | -34.1188 | -7.7173
4.2852 | -34.119 | -7.7279
4.2913 | -34.1191 | -7.7381
4.2974 | -34.1194 | -7.7488
4.3035 | -34.1196 | -7.7595
4.3096 | -34.1198 | -7.7702
4.3157 | -34.1201 | -7.7804
4.3218 | -34.1203 | -7.7908
4.3279 | -34.1205 | -7.8012
4.334 | -34.1207 | -7.8116
4.3401 | -34.1209 | -7.8216
4.3462 | -34.1212 | -7.8321
4.3523 | -34.1214 | -7.8421
4.3584 | -34.1216 | -7.8521
4.3645 | -34.1218 | -7.8624
4.3707 | -34.122 | -7.8725
4.3768 | -34.1222 | -7.8828
4.3829 | -34.1225 | -7.8935
4.389 | -34.1227 | -7.9039
4.3951 | -34.1229 | -7.914
4.4012 | -34.1231 | -7.9247
4.4073 | -34.1232 | -7.9349
4.4134 | -34.1235 | -7.9451
4.4195 | -34.1237 | -7.9554
4.4256 | -34.124 | -7.9655
4.4317 | -34.1242 | -7.9758
4.4378 | -34.1244 | -7.9857
4.4439 | -34.1246 | -7.9958
4.45 | -34.1247 | -8.006
4.4561 | -34.1249 | -8.0163
4.4622 | -34.1251 | -8.0267
4.4683 | -34.1253 | -8.0369
4.4744 | -34.1255 | -8.0475
4.4805 | -34.1257 | -8.0574
4.4866 | -34.1259 | -8.0676
4.4927 | -34.1261 | -8.0782
4.4988 | -34.1263 | -8.0884
4.5049 | -34.1266 | -8.0982
4.511 | -34.1268 | -8.1082
4.5172 | -34.127 | -8.1184
4.5233 | -34.1272 | -8.129
4.5294 | -34.1274 | -8.1391
4.5355 | -34.1276 | -8.1494
4.5416 | -34.1278 | -8.1596
4.5477 | -34.128 | -8.1698
4.5538 | -34.1282 | -8.1796
4.5599 | -34.1284 | -8.1898
4.566 | -34.1286 | -8.2
4.5721 | -34.1288 | -8.2102
4.5782 | -34.1291 | -8.2207
4.5843 | -34.1293 | -8.231
4.5904 | -34.1295 | -8.2413
4.5965 | -34.1298 | -8.2514
4.6026 | -34.13 | -8.2621
4.6087 | -34.1302 | -8.2722
4.6148 | -34.1304 | -8.2828
4.6209 | -34.1306 | -8.2929
4.627 | -34.1308 | -8.3034
4.6331 | -34.131 | -8.3135
4.6392 | -34.1312 | -8.3238
4.6453 | -34.1314 | -8.334
4.6514 | -34.1317 | -8.3443
4.6576 | -34.1318 | -8.3549
4.6637 | -34.132 | -8.3648
4.6698 | -34.1322 | -8.3749
4.6759 | -34.1324 | -8.3855
4.682 | -34.1326 | -8.3956
4.6881 | -34.1327 | -8.4058
4.6942 | -34.133 | -8.416
4.7003 | -34.1332 | -8.4263
4.7064 | -34.1334 | -8.4364
4.7125 | -34.1336 | -8.4469
4.7186 | -34.1338 | -8.4569
4.7247 | -34.1341 | -8.4672
4.7308 | -34.1343 | -8.4774
4.7369 | -34.1345 | -8.4878
4.743 | -34.1347 | -8.4984
4.7491 | -34.1349 | -8.5089
4.7552 | -34.1351 | -8.5193
4.7613 | -34.1353 | -8.5297
4.7674 | -34.1355 | -8.5398
4.7735 | -34.1357 | -8.5506
4.7796 | -34.136 | -8.5611
4.7857 | -34.1362 | -8.5715
4.7918 | -34.1364 | -8.5816
4.7979 | -34.1367 | -8.5918
4.8041 | -34.1369 | -8.6021
4.8102 | -34.1371 | -8.6121
4.8163 | -34.1373 | -8.6224
4.8224 | -34.1375 | -8.6327
4.8285 | -34.1378 | -8.6427
4.8346 | -34.138 | -8.6528
4.8407 | -34.1382 | -8.6632
4.8468 | -34.1385 | -8.6737
4.8529 | -34.1387 | -8.6843
4.859 | -34.1389 | -8.695
4.8651 | -34.1391 | -8.705
4.8712 | -34.1394 | -8.7151
4.8773 | -34.1395 | -8.7255
4.8834 | -34.1398 | -8.7358
4.8895 | -34.14 | -8.7459
4.8956 | -34.1402 | -8.7562
4.9017 | -34.1404 | -8.7664
4.9078 | -34.1406 | -8.777
4.9139 | -34.1408 | -8.7869
4.92 | -34.1411 | -8.7969
4.9261 | -34.1414 | -8.8068
4.9322 | -34.1416 | -8.8172
4.9383 | -34.1419 | -8.8274
4.9445 | -34.1421 | -8.8378
4.9506 | -34.1423 | -8.8479
4.9567 | -34.1426 | -8.8579
4.9628 | -34.1428 | -8.8682
4.9689 | -34.143 | -8.8788
4.975 | -34.1432 | -8.8894
4.9811 | -34.1435 | -8.8994
4.9872 | -34.1437 | -8.9096
4.9933 | -34.144 | -8.9202
4.9994 | -34.1443 | -8.9309
5.0055 | -34.1445 | -8.9413
5.0116 | -34.1448 | -8.952
5.0177 | -34.1451 | -8.9626
5.0238 | -34.1454 | -8.9732
5.0299 | -34.1457 | -8.9839
5.036 | -34.1459 | -8.9942
5.0421 | -34.1461 | -9.0044
5.0482 | -34.1464 | -9.0148
5.0543 | -34.1467 | -9.0249
5.0604 | -34.147 | -9.0352
5.0665 | -34.1472 | -9.0453
5.0726 | -34.1475 | -9.0559
5.0787 | -34.1477 | -9.0665
5.0848 | -34.148 | -9.077
5.091 | -34.1483 | -9.0871
5.0971 | -34.1485 | -9.0971
5.1032 | -34.1488 | -9.1078
5.1093 | -34.1491 | -9.1182
5.1154 | -34.1493 | -9.1283
5.1215 | -34.1495 | -9.1384
5.1276 | -34.1498 | -9.1486
5.1337 | -34.1501 | -9.1586
5.1398 | -34.1504 | -9.1687
5.1459 | -34.1506 | -9.1793
5.152 | -34.1509 | -9.1899
5.1581 | -34.1511 | -9.2004
5.1642 | -34.1514 | -9.2108
5.1703 | -34.1517 | -9.2212
5.1764 | -34.152 | -9.2314
5.1825 | -34.1523 | -9.2415
5.1886 | -34.1526 | -9.2515
5.1947 | -34.1528 | -9.2619
5.2008 | -34.1531 | -9.2724
5.2069 | -34.1534 | -9.283
5.213 | -34.1536 | -9.2931
5.2191 | -34.1539 | -9.3035
5.2252 | -34.1542 | -9.314
5.2314 | -34.1544 | -9.3245
5.2375 | -34.1547 | -9.3346
5.2436 | -34.155 | -9.3451
5.2497 | -34.1553 | -9.3552
5.2558 | -34.1556 | -9.3659
5.2619 | -34.1558 | -9.3765
5.268 | -34.1561 | -9.3871
5.2741 | -34.1564 | -9.3975
5.2802 | -34.1567 | -9.4083
5.2863 | -34.157 | -9.4186
5.2924 | -34.1573 | -9.4292
5.2985 | -34.1576 | -9.4396
5.3046 | -34.1579 | -9.4501
5.3107 | -34.1582 | -9.4607
5.3168 | -34.1585 | -9.4712
5.3229 | -34.1588 | -9.4817
5.329 | -34.159 | -9.4921
5.3351 | -34.1593 | -9.5022
5.3412 | -34.1596 | -9.5126
5.3473 | -34.1599 | -9.5231
5.3534 | -34.1602 | -9.5336
5.3595 | -34.1605 | -9.5441
5.3656 | -34.1608 | -9.5545
5.3717 | -34.1611 | -9.5653
5.3779 | -34.1614 | -9.576
5.384 | -34.1617 | -9.5865
5.3901 | -34.162 | -9.5972
5.3962 | -34.1622 | -9.6074
5.4023 | -34.1626 | -9.6181
5.4084 | -34.1629 | -9.6285
5.4145 | -34.1632 | -9.6387
5.4206 | -34.1635 | -9.6491
5.4267 | -34.1638 | -9.6599
5.4328 | -34.1641 | -9.6705
5.4389 | -34.1643 | -9.6807
5.445 | -34.1646 | -9.6915
5.4511 | -34.1649 | -9.7017
5.4572 | -34.1652 | -9.7123
5.4633 | -34.1655 | -9.7227
5.4694 | -34.1658 | -9.7332
5.4755 | -34.1661 | -9.7434
5.4816 | -34.1664 | -9.7539
5.4877 | -34.1667 | -9.7643
5.4938 | -34.167 | -9.7745
5.4999 | -34.1673 | -9.7848
5.506 | -34.1676 | -9.7953
5.5121 | -34.1679 | -9.8056
5.5183 | -34.1682 | -9.8159
5.5244 | -34.1685 | -9.8266
5.5305 | -34.1688 | -9.837
5.5366 | -34.1691 | -9.8472
5.5427 | -34.1693 | -9.8578
5.5488 | -34.1696 | -9.8683
5.5549 | -34.1699 | -9.8788
5.561 | -34.1702 | -9.8895
5.5671 | -34.1705 | -9.8998
5.5732 | -34.1708 | -9.9105
5.5793 | -34.1712 | -9.9213
5.5854 | -34.1715 | -9.9317
5.5915 | -34.1718 | -9.9425
5.5976 | -34.1721 | -9.953
5.6037 | -34.1723 | -9.9636
5.6098 | -34.1726 | -9.974
5.6159 | -34.1729 | -9.9843
5.622 | -34.1732 | -9.9948
5.6281 | -34.1735 | -10.0054
5.6342 | -34.1738 | -10.0162
5.6403 | -34.1741 | -10.0267
5.6464 | -34.1744 | -10.0374
5.6525 | -34.1747 | -10.0479
5.6586 | -34.175 | -10.0582
5.6648 | -34.1754 | -10.0687
5.6709 | -34.1757 | -10.0793
5.677 | -34.176 | -10.0899
5.6831 | -34.1763 | -10.1004
5.6892 | -34.1766 | -10.1111
5.6953 | -34.1769 | -10.1216
5.7014 | -34.1772 | -10.1322
5.7075 | -34.1775 | -10.1427
5.7136 | -34.1778 | -10.1531
5.7197 | -34.178 | -10.1636
5.7258 | -34.1783 | -10.1739
5.7319 | -34.1787 | -10.1845
5.738 | -34.179 | -10.1949
5.7441 | -34.1793 | -10.2054
5.7502 | -34.1796 | -10.2156
5.7563 | -34.1799 | -10.2261
5.7624 | -34.1803 | -10.2368
5.7685 | -34.1806 | -10.2473
5.7746 | -34.1808 | -10.2581
5.7807 | -34.1811 | -10.2686
5.7868 | -34.1814 | -10.2793
5.7929 | -34.1817 | -10.2897
5.799 | -34.182 | -10.3001
5.8052 | -34.1823 | -10.3104
5.8113 | -34.1826 | -10.3209
5.8174 | -34.1828 | -10.3312
5.8235 | -34.1832 | -10.3416
5.8296 | -34.1834 | -10.3523
5.8357 | -34.1837 | -10.3628
5.8418 | -34.184 | -10.3734
5.8479 | -34.1843 | -10.3839
5.854 | -34.1846 | -10.3945
5.8601 | -34.1849 | -10.4049
5.8662 | -34.1852 | -10.4155
5.8723 | -34.1855 | -10.4258
5.8784 | -34.1857 | -10.4362
5.8845 | -34.186 | -10.4466
5.8906 | -34.1863 | -10.4571
5.8967 | -34.1866 | -10.4674
5.9028 | -34.1869 | -10.4777
5.9089 | -34.1873 | -10.488
5.915 | -34.1876 | -10.4983
5.9211 | -34.1878 | -10.5087
5.9272 | -34.1881 | -10.5189
5.9333 | -34.1884 | -10.5292
5.9394 | -34.1887 | -10.5397
5.9455 | -34.1891 | -10.5499
5.9517 | -34.1894 | -10.5602
5.9578 | -34.1897 | -10.5707
5.9639 | -34.19 | -10.5808
5.97 | -34.1904 | -10.5913
5.9761 | -34.1907 | -10.6016
5.9822 | -34.191 | -10.6121
5.9883 | -34.1913 | -10.6225
5.9944 | -34.1916 | -10.633
6.0005 | -34.1919 | -10.6436
6.0066 | -34.1922 | -10.6538
6.0127 | -34.1926 | -10.664
6.0188 | -34.1928 | -10.6747
6.0249 | -34.1931 | -10.6852
6.031 | -34.1934 | -10.6954
6.0371 | -34.1937 | -10.7058
6.0432 | -34.194 | -10.7161
6.0493 | -34.1943 | -10.7263
6.0554 | -34.1947 | -10.7365
6.0615 | -34.195 | -10.7469
6.0676 | -34.1953 | -10.7575
6.0737 | -34.1956 | -10.768
6.0798 | -34.1959 | -10.7784
6.0859 | -34.1963 | -10.7887
6.0921 | -34.1965 | -10.7988
6.0982 | -34.1968 | -10.8092
6.1043 | -34.1972 | -10.8194
6.1104 | -34.1975 | -10.83
6.1165 | -34.1978 | -10.8399
6.1226 | -34.1982 | -10.8502
6.1287 | -34.1984 | -10.8605
6.1348 | -34.1988 | -10.8708
6.1409 | -34.1991 | -10.8811
6.147 | -34.1994 | -10.8914
6.1531 | -34.1997 | -10.9017
6.1592 | -34.2 | -10.9119
6.1653 | -34.2003 | -10.9221
6.1714 | -34.2006 | -10.9324
6.1775 | -34.2009 | -10.9429
6.1836 | -34.2012 | -10.9532
6.1897 | -34.2015 | -10.9635
6.1958 | -34.2018 | -10.9741
6.2019 | -34.2022 | -10.9845
6.208 | -34.2025 | -10.9948
6.2141 | -34.2028 | -11.0051
6.2202 | -34.2031 | -11.0155
6.2263 | -34.2034 | -11.0257
6.2325 | -34.2037 | -11.0357
6.2386 | -34.204 | -11.0457
6.2447 | -34.2044 | -11.0563
6.2508 | -34.2047 | -11.0665
6.2569 | -34.205 | -11.0768
6.263 | -34.2053 | -11.087
6.2691 | -34.2056 | -11.0972
6.2752 | -34.2059 | -11.1076
6.2813 | -34.2062 | -11.1179
6.2874 | -34.2065 | -11.128
6.2935 | -34.2068 | -11.1382
6.2996 | -34.2071 | -11.1485
6.3057 | -34.2074 | -11.1588
6.3118 | -34.2077 | -11.1692
6.3179 | -34.208 | -11.1794
6.324 | -34.2083 | -11.1896
6.3301 | -34.2086 | -11.1998
6.3362 | -34.209 | -11.21
6.3423 | -34.2093 | -11.2204
6.3484 | -34.2096 | -11.2307
6.3545 | -34.2099 | -11.2406
6.3606 | -34.2102 | -11.2506
6.3667 | -34.2104 | -11.2607
6.3728 | -34.2108 | -11.271
6.379 | -34.2111 | -11.2812
6.3851 | -34.2114 | -11.2912
6.3912 | -34.2117 | -11.3014
6.3973 | -34.212 | -11.3116
6.4034 | -34.2123 | -11.3216
6.4095 | -34.2126 | -11.3318
6.4156 | -34.2129 | -11.3421
6.4217 | -34.2132 | -11.3522
6.4278 | -34.2135 | -11.3625
6.4339 | -34.2138 | -11.3726
6.44 | -34.2141 | -11.3827
6.4461 | -34.2144 | -11.3931
6.4522 | -34.2147 | -11.4031
6.4583 | -34.215 | -11.4134
6.4644 | -34.2153 | -11.4238
6.4705 | -34.2156 | -11.4336
6.4766 | -34.2159 | -11.4438
6.4827 | -34.2162 | -11.454
6.4888 | -34.2165 | -11.464
6.4949 | -34.2168 | -11.4739
6.501 | -34.2171 | -11.4842
6.5071 | -34.2174 | -11.4941
6.5132 | -34.2177 | -11.5043
6.5194 | -34.218 | -11.5145
6.5255 | -34.2183 | -11.5244
6.5316 | -34.2186 | -11.5347
6.5377 | -34.2189 | -11.5448
6.5438 | -34.2191 | -11.5547
6.5499 | -34.2195 | -11.5647
6.556 | -34.2198 | -11.5748
6.5621 | -34.2201 | -11.5852
6.5682 | -34.2204 | -11.5954
6.5743 | -34.2207 | -11.6054
6.5804 | -34.221 | -11.6157
6.5865 | -34.2213 | -11.6258
6.5926 | -34.2216 | -11.6358
6.5987 | -34.2219 | -11.6457
6.6048 | -34.2222 | -11.6558
6.6109 | -34.2225 | -11.6659
6.617 | -34.2228 | -11.6762
6.6231 | -34.2231 | -11.6862
6.6292 | -34.2234 | -11.6964
6.6353 | -34.2237 | -11.7068
6.6414 | -34.2241 | -11.7168
6.6475 | -34.2244 | -11.7268
6.6536 | -34.2246 | -11.7369
6.6597 | -34.225 | -11.7467
6.6659 | -34.2252 | -11.7568
6.672 | -34.2256 | -11.7668
6.6781 | -34.2259 | -11.7769
6.6842 | -34.2262 | -11.7868
6.6903 | -34.2265 | -11.7969
6.6964 | -34.2269 | -11.807
6.7025 | -34.2271 | -11.8168
6.7086 | -34.2274 | -11.8269
6.7147 | -34.2277 | -11.837
6.7208 | -34.228 | -11.8472
6.7269 | -34.2283 | -11.8572
6.733 | -34.2286 | -11.867
6.7391 | -34.2289 | -11.8771
6.7452 | -34.2293 | -11.8869
6.7513 | -34.2295 | -11.8967
6.7574 | -34.2299 | -11.9068
6.7635 | -34.2302 | -11.9167
6.7696 | -34.2305 | -11.9265
6.7757 | -34.2308 | -11.9366
6.7818 | -34.2311 | -11.9466
6.7879 | -34.2315 | -11.9565
6.794 | -34.2318 | -11.9663
6.8001 | -34.232 | -11.9762
6.8063 | -34.2323 | -11.9862
6.8124 | -34.2326 | -11.9964
6.8185 | -34.2329 | -12.0066
6.8246 | -34.2332 | -12.0164
6.8307 | -34.2335 | -12.0262
6.8368 | -34.2338 | -12.0365
6.8429 | -34.2342 | -12.0465
6.849 | -34.2345 | -12.0565
6.8551 | -34.2348 | -12.0664
6.8612 | -34.2351 | -12.0763
6.8673 | -34.2354 | -12.0863
6.8734 | -34.2357 | -12.0964
6.8795 | -34.236 | -12.1064
6.8856 | -34.2364 | -12.1164
6.8917 | -34.2367 | -12.1263
6.8978 | -34.237 | -12.136
6.9039 | -34.2373 | -12.1457
6.91 | -34.2376 | -12.1557
6.9161 | -34.2379 | -12.1654
6.9222 | -34.2382 | -12.1754
6.9283 | -34.2386 | -12.185
6.9344 | -34.2389 | -12.195
6.9405 | -34.2393 | -12.2049
6.9466 | -34.2396 | -12.2151
6.9528 | -34.2399 | -12.2252
6.9589 | -34.2402 | -12.2353
6.965 | -34.2406 | -12.2453
6.9711 | -34.2409 | -12.2554
6.9772 | -34.2413 | -12.2653
6.9833 | -34.2416 | -12.2752
6.9894 | -34.2419 | -12.2854
6.9955 | -34.2422 | -12.2951
7.0016 | -34.2425 | -12.305
7.0077 | -34.2429 | -12.3149
7.0138 | -34.2432 | -12.3248
7.0199 | -34.2435 | -12.3346
7.026 | -34.2438 | -12.3445
7.0321 | -34.2441 | -12.3542
7.0382 | -34.2445 | -12.3638
7.0443 | -34.2448 | -12.3739
7.0504 | -34.2451 | -12.3839
7.0565 | -34.2455 | -12.3938
7.0626 | -34.2458 | -12.4035
7.0687 | -34.2461 | -12.413
7.0748 | -34.2465 | -12.4226
7.0809 | -34.2468 | -12.4325
7.087 | -34.2471 | -12.4425
7.0932 | -34.2474 | -12.4524
7.0993 | -34.2477 | -12.4622
7.1054 | -34.2481 | -12.472
7.1115 | -34.2484 | -12.4816
7.1176 | -34.2487 | -12.4914
7.1237 | -34.249 | -12.5012
7.1298 | -34.2494 | -12.511
7.1359 | -34.2497 | -12.5208
7.142 | -34.25 | -12.5304
7.1481 | -34.2503 | -12.5402
7.1542 | -34.2506 | -12.5501
7.1603 | -34.251 | -12.5599
7.1664 | -34.2513 | -12.5697
7.1725 | -34.2517 | -12.5796
7.1786 | -34.252 | -12.5894
7.1847 | -34.2523 | -12.5993
7.1908 | -34.2527 | -12.6093
7.1969 | -34.2529 | -12.6193
7.203 | -34.2533 | -12.629
7.2091 | -34.2536 | -12.6387
7.2152 | -34.254 | -12.6486
7.2213 | -34.2543 | -12.6582
7.2274 | -34.2546 | -12.668
7.2335 | -34.255 | -12.6778
7.2397 | -34.2553 | -12.6875
7.2458 | -34.2557 | -12.6971
7.2519 | -34.256 | -12.7069
7.258 | -34.2563 | -12.7169
7.2641 | -34.2567 | -12.7267
7.2702 | -34.257 | -12.7367
7.2763 | -34.2573 | -12.7466
7.2824 | -34.2577 | -12.7565
7.2885 | -34.258 | -12.7664
7.2946 | -34.2583 | -12.7762
7.3007 | -34.2586 | -12.7862
7.3068 | -34.259 | -12.7962
7.3129 | -34.2593 | -12.8062
7.319 | -34.2597 | -12.816
7.3251 | -34.26 | -12.8257
7.3312 | -34.2603 | -12.8355
7.3373 | -34.2607 | -12.8451
7.3434 | -34.261 | -12.8548
7.3495 | -34.2614 | -12.8645
7.3556 | -34.2617 | -12.8744
7.3617 | -34.2621 | -12.8844
7.3678 | -34.2624 | -12.8942
7.3739 | -34.2627 | -12.9042
7.3801 | -34.263 | -12.9142
7.3862 | -34.2634 | -12.924
7.3923 | -34.2638 | -12.934
7.3984 | -34.2641 | -12.9435
7.4045 | -34.2645 | -12.9533
7.4106 | -34.2648 | -12.9631
7.4167 | -34.2651 | -12.973
7.4228 | -34.2655 | -12.9829
7.4289 | -34.2658 | -12.9927
7.435 | -34.2662 | -13.0023
7.4411 | -34.2665 | -13.0122
7.4472 | -34.2669 | -13.0221
7.4533 | -34.2672 | -13.0318
7.4594 | -34.2676 | -13.0418
7.4655 | -34.2679 | -13.0515
7.4716 | -34.2683 | -13.0614
7.4777 | -34.2686 | -13.0712
7.4838 | -34.2689 | -13.0809
7.4899 | -34.2693 | -13.0907
7.496 | -34.2696 | -13.1004
7.5021 | -34.2699 | -13.1101
7.5082 | -34.2703 | -13.12
7.5143 | -34.2706 | -13.1298
7.5204 | -34.271 | -13.1398
7.5266 | -34.2713 | -13.1495
7.5327 | -34.2716 | -13.1592
7.5388 | -34.272 | -13.1693
7.5449 | -34.2723 | -13.1793
7.551 | -34.2727 | -13.189
7.5571 | -34.273 | -13.1989
7.5632 | -34.2734 | -13.2088
7.5693 | -34.2738 | -13.2185
7.5754 | -34.2741 | -13.2283
7.5815 | -34.2744 | -13.2381
7.5876 | -34.2748 | -13.2482
7.5937 | -34.2752 | -13.2582
7.5998 | -34.2756 | -13.2679
7.6059 | -34.2759 | -13.2779
7.612 | -34.2762 | -13.2877
7.6181 | -34.2766 | -13.2975
7.6242 | -34.2769 | -13.3076
7.6303 | -34.2773 | -13.3177
7.6364 | -34.2776 | -13.3276
7.6425 | -34.278 | -13.3374
7.6486 | -34.2783 | -13.3473
7.6547 | -34.2786 | -13.357
7.6608 | -34.279 | -13.3668
7.667 | -34.2793 | -13.3768
7.6731 | -34.2797 | -13.3868
7.6792 | -34.2801 | -13.3968
7.6853 | -34.2804 | -13.4068
7.6914 | -34.2808 | -13.4167
7.6975 | -34.2811 | -13.4267
7.7036 | -34.2815 | -13.4367
7.7097 | -34.2818 | -13.4467
7.7158 | -34.2822 | -13.4567
7.7219 | -34.2826 | -13.4664
7.728 | -34.2829 | -13.4764
7.7341 | -34.2832 | -13.4864
7.7402 | -34.2836 | -13.4963
7.7463 | -34.284 | -13.5063
7.7524 | -34.2843 | -13.5164
7.7585 | -34.2846 | -13.5264
7.7646 | -34.285 | -13.5362
7.7707 | -34.2853 | -13.5464
7.7768 | -34.2857 | -13.5564
7.7829 | -34.286 | -13.5666
7.789 | -34.2863 | -13.5767
7.7951 | -34.2867 | -13.5866
7.8012 | -34.287 | -13.5965
7.8073 | -34.2875 | -13.6067
7.8135 | -34.2878 | -13.6166
7.8196 | -34.2882 | -13.6268
7.8257 | -34.2885 | -13.637
7.8318 | -34.2888 | -13.647
7.8379 | -34.2892 | -13.6568
7.844 | -34.2896 | -13.6667
7.8501 | -34.2899 | -13.6768
7.8562 | -34.2903 | -13.687
7.8623 | -34.2907 | -13.6968
7.8684 | -34.291 | -13.7069
7.8745 | -34.2913 | -13.7169
7.8806 | -34.2917 | -13.727
7.8867 | -34.292 | -13.7368
7.8928 | -34.2924 | -13.7465
7.8989 | -34.2928 | -13.7565
7.905 | -34.2931 | -13.7666
7.9111 | -34.2934 | -13.7766
7.9172 | -34.2938 | -13.7865
7.9233 | -34.2941 | -13.7966
7.9294 | -34.2945 | -13.8066
7.9355 | -34.2948 | -13.8164
7.9416 | -34.2952 | -13.8264
7.9477 | -34.2955 | -13.8363
7.9539 | -34.2959 | -13.8462
7.96 | -34.2962 | -13.8563
7.9661 | -34.2966 | -13.8663
7.9722 | -34.297 | -13.8762
7.9783 | -34.2973 | -13.8864
7.9844 | -34.2977 | -13.8963
7.9905 | -34.2981 | -13.9065
7.9966 | -34.2984 | -13.9166
8.0027 | -34.2988 | -13.9266
8.0088 | -34.2991 | -13.9366
8.0149 | -34.2994 | -13.9466
8.021 | -34.2998 | -13.9566
8.0271 | -34.3001 | -13.9669
8.0332 | -34.3005 | -13.9768
8.0393 | -34.3009 | -13.987
8.0454 | -34.3012 | -13.9973
8.0515 | -34.3016 | -14.0073
8.0576 | -34.302 | -14.0175
8.0637 | -34.3023 | -14.0275
8.0698 | -34.3026 | -14.0373
8.0759 | -34.3029 | -14.0473
8.082 | -34.3033 | -14.0572
8.0881 | -34.3036 | -14.0673
8.0942 | -34.304 | -14.0774
8.1004 | -34.3044 | -14.0875
8.1065 | -34.3048 | -14.0973
8.1126 | -34.3051 | -14.1073
8.1187 | -34.3054 | -14.1174
8.1248 | -34.3058 | -14.1272
8.1309 | -34.3061 | -14.1371
8.137 | -34.3065 | -14.1469
8.1431 | -34.3069 | -14.1571
8.1492 | -34.3072 | -14.1671
8.1553 | -34.3076 | -14.1771
8.1614 | -34.3079 | -14.1873
8.1675 | -34.3083 | -14.1973
8.1736 | -34.3086 | -14.2075
8.1797 | -34.3089 | -14.2176
8.1858 | -34.3093 | -14.2277
8.1919 | -34.3096 | -14.2377
8.198 | -34.31 | -14.2475
8.2041 | -34.3104 | -14.2576
8.2102 | -34.3107 | -14.2676
8.2163 | -34.3111 | -14.2774
8.2224 | -34.3114 | -14.2874
8.2285 | -34.3118 | -14.2972
8.2346 | -34.3121 | -14.3071
8.2408 | -34.3125 | -14.3169
8.2469 | -34.3129 | -14.327
8.253 | -34.3132 | -14.3369
8.2591 | -34.3136 | -14.347
8.2652 | -34.314 | -14.357
8.2713 | -34.3143 | -14.367
8.2774 | -34.3147 | -14.3772
8.2835 | -34.315 | -14.3874
8.2896 | -34.3154 | -14.3973
8.2957 | -34.3158 | -14.4074
8.3018 | -34.3161 | -14.4174
8.3079 | -34.3165 | -14.4272
8.314 | -34.3169 | -14.4372
8.3201 | -34.3172 | -14.4471
8.3262 | -34.3176 | -14.4572
8.3323 | -34.3179 | -14.4672
8.3384 | -34.3183 | -14.4773
8.3445 | -34.3186 | -14.4872
8.3506 | -34.319 | -14.4973
8.3567 | -34.3194 | -14.5074
8.3628 | -34.3197 | -14.5173
8.3689 | -34.3201 | -14.5271
8.375 | -34.3204 | -14.5372
8.3812 | -34.3208 | -14.5471
8.3873 | -34.3212 | -14.5571
8.3934 | -34.3215 | -14.5672
8.3995 | -34.3218 | -14.577
8.4056 | -34.3222 | -14.587
8.4117 | -34.3226 | -14.5971
8.4178 | -34.3229 | -14.6071
8.4239 | -34.3233 | -14.6169
8.43 | -34.3236 | -14.6268
8.4361 | -34.3239 | -14.6369
8.4422 | -34.3243 | -14.6468
8.4483 | -34.3247 | -14.6566
8.4544 | -34.325 | -14.6666
8.4605 | -34.3254 | -14.6765
8.4666 | -34.3258 | -14.6867
8.4727 | -34.3261 | -14.6967
8.4788 | -34.3265 | -14.7066
8.4849 | -34.3268 | -14.7166
8.491 | -34.3272 | -14.7266
8.4971 | -34.3275 | -14.7363
8.5032 | -34.3279 | -14.7461
8.5093 | -34.3282 | -14.7559
8.5154 | -34.3286 | -14.766
8.5215 | -34.329 | -14.7758
8.5277 | -34.3293 | -14.7858
8.5338 | -34.3297 | -14.7956
8.5399 | -34.33 | -14.8056
8.546 | -34.3304 | -14.8153
8.5521 | -34.3308 | -14.8253
8.5582 | -34.3311 | -14.8352
8.5643 | -34.3315 | -14.8448
8.5704 | -34.3318 | -14.8545
8.5765 | -34.3322 | -14.8646
8.5826 | -34.3326 | -14.8742
8.5887 | -34.333 | -14.884
8.5948 | -34.3333 | -14.894
8.6009 | -34.3337 | -14.9037
8.607 | -34.334 | -14.9138
8.6131 | -34.3344 | -14.9238
8.6192 | -34.3348 | -14.9335
8.6253 | -34.3352 | -14.9433
8.6314 | -34.3355 | -14.9531
8.6375 | -34.3358 | -14.9631
8.6436 | -34.3362 | -14.9729
8.6497 | -34.3366 | -14.9827
8.6558 | -34.3369 | -14.9926
8.6619 | -34.3372 | -15.0025
8.6681 | -34.3376 | -15.0125
8.6742 | -34.338 | -15.0224
8.6803 | -34.3383 | -15.0322
8.6864 | -34.3387 | -15.0418
8.6925 | -34.3391 | -15.0515
8.6986 | -34.3394 | -15.0615
8.7047 | -34.3398 | -15.0716
8.7108 | -34.3401 | -15.0814
8.7169 | -34.3405 | -15.0913
8.723 | -34.3408 | -15.101
8.7291 | -34.3412 | -15.1107
8.7352 | -34.3416 | -15.1208
8.7413 | -34.3419 | -15.1307
8.7474 | -34.3423 | -15.1404
8.7535 | -34.3426 | -15.1502
8.7596 | -34.343 | -15.1601
8.7657 | -34.3434 | -15.1699
8.7718 | -34.3437 | -15.1798
8.7779 | -34.3441 | -15.1897
8.784 | -34.3444 | -15.1996
8.7901 | -34.3448 | -15.2094
8.7962 | -34.3452 | -15.219
8.8023 | -34.3455 | -15.2288
8.8084 | -34.3459 | -15.2386
8.8146 | -34.3463 | -15.2484
8.8207 | -34.3466 | -15.2582
8.8268 | -34.347 | -15.2677
8.8329 | -34.3473 | -15.2775
8.839 | -34.3477 | -15.2874
8.8451 | -34.3481 | -15.2971
8.8512 | -34.3485 | -15.3068
8.8573 | -34.3488 | -15.3166
8.8634 | -34.3492 | -15.3262
8.8695 | -34.3496 | -15.3362
8.8756 | -34.3499 | -15.346
8.8817 | -34.3503 | -15.3556
8.8878 | -34.3506 | -15.3655
8.8939 | -34.351 | -15.3753
8.9 | -34.3514 | -15.3851
8.9061 | -34.3517 | -15.3949
8.9122 | -34.3521 | -15.4047
8.9183 | -34.3525 | -15.4145
8.9244 | -34.3529 | -15.4244
8.9305 | -34.3533 | -15.434
8.9366 | -34.3536 | -15.4437
8.9427 | -34.354 | -15.4536
8.9488 | -34.3544 | -15.4632
8.955 | -34.3547 | -15.473
8.9611 | -34.3551 | -15.4829
8.9672 | -34.3555 | -15.4926
8.9733 | -34.3559 | -15.5022
8.9794 | -34.3563 | -15.5118
8.9855 | -34.3567 | -15.5217
8.9916 | -34.357 | -15.5316
8.9977 | -34.3574 | -15.5413
9.0038 | -34.3578 | -15.5509
9.0099 | -34.3582 | -15.5607
9.016 | -34.3586 | -15.5705
9.0221 | -34.359 | -15.58
9.0282 | -34.3593 | -15.5897
9.0343 | -34.3597 | -15.5995
9.0404 | -34.3601 | -15.6093
9.0465 | -34.3605 | -15.6192
9.0526 | -34.3609 | -15.6288
9.0587 | -34.3613 | -15.6387
9.0648 | -34.3617 | -15.6482
9.0709 | -34.3621 | -15.658
9.077 | -34.3625 | -15.6678
9.0831 | -34.3629 | -15.6777
9.0892 | -34.3633 | -15.6875
9.0953 | -34.3637 | -15.6974
9.1015 | -34.3641 | -15.7071
9.1076 | -34.3645 | -15.7168
9.1137 | -34.3648 | -15.7265
9.1198 | -34.3652 | -15.7362
9.1259 | -34.3656 | -15.746
9.132 | -34.366 | -15.7558
9.1381 | -34.3664 | -15.7654
9.1442 | -34.3668 | -15.7748
9.1503 | -34.3672 | -15.7847
9.1564 | -34.3676 | -15.7943
9.1625 | -34.368 | -15.8041
9.1686 | -34.3684 | -15.8138
9.1747 | -34.3688 | -15.8235
9.1808 | -34.3692 | -15.833
9.1869 | -34.3696 | -15.8427
9.193 | -34.37 | -15.8524
9.1991 | -34.3703 | -15.8622
9.2052 | -34.3708 | -15.8719
9.2113 | -34.3712 | -15.8816
9.2174 | -34.3716 | -15.8913
9.2235 | -34.372 | -15.901
9.2296 | -34.3724 | -15.9107
9.2357 | -34.3728 | -15.9205
9.2419 | -34.3732 | -15.9302
9.248 | -34.3737 | -15.9399
9.2541 | -34.3741 | -15.9497
9.2602 | -34.3745 | -15.9595
9.2663 | -34.3749 | -15.9689
9.2724 | -34.3753 | -15.9786
9.2785 | -34.3757 | -15.9883
9.2846 | -34.3761 | -15.9979
9.2907 | -34.3765 | -16.0076
9.2968 | -34.3769 | -16.0174
9.3029 | -34.3773 | -16.0272
9.309 | -34.3777 | -16.0369
9.3151 | -34.3781 | -16.0465
9.3212 | -34.3785 | -16.056
9.3273 | -34.3789 | -16.0656
9.3334 | -34.3793 | -16.0752
9.3395 | -34.3797 | -16.0847
9.3456 | -34.3801 | -16.0943
9.3517 | -34.3804 | -16.104
9.3578 | -34.3809 | -16.1137
9.3639 | -34.3813 | -16.1232
9.37 | -34.3817 | -16.1329
9.3761 | -34.3821 | -16.1428
9.3822 | -34.3825 | -16.1524
9.3884 | -34.3829 | -16.1618
9.3945 | -34.3834 | -16.1715
9.4006 | -34.3838 | -16.1812
9.4067 | -34.3842 | -16.191
9.4128 | -34.3846 | -16.2007
9.4189 | -34.385 | -16.2106
9.425 | -34.3854 | -16.2201
9.4311 | -34.3859 | -16.2297
9.4372 | -34.3863 | -16.2394
9.4433 | -34.3867 | -16.249
9.4494 | -34.3871 | -16.2585
9.4555 | -34.3876 | -16.2683
9.4616 | -34.388 | -16.2779
9.4677 | -34.3884 | -16.2876
9.4738 | -34.3888 | -16.2972
9.4799 | -34.3892 | -16.3069
9.486 | -34.3896 | -16.3167
9.4921 | -34.39 | -16.3263
9.4982 | -34.3905 | -16.3358
9.5043 | -34.3909 | -16.3454
9.5104 | -34.3913 | -16.3548
9.5165 | -34.3917 | -16.3645
9.5226 | -34.3921 | -16.3743
9.5288 | -34.3925 | -16.3838
9.5349 | -34.3929 | -16.3934
9.541 | -34.3933 | -16.4029
9.5471 | -34.3937 | -16.4125
9.5532 | -34.3941 | -16.4222
9.5593 | -34.3946 | -16.4319
9.5654 | -34.395 | -16.4415
9.5715 | -34.3954 | -16.4511
9.5776 | -34.3958 | -16.4609
9.5837 | -34.3962 | -16.4705
9.5898 | -34.3966 | -16.4801
9.5959 | -34.397 | -16.4897
9.602 | -34.3974 | -16.4991
9.6081 | -34.3978 | -16.5087
9.6142 | -34.3983 | -16.5183
9.6203 | -34.3987 | -16.5279
9.6264 | -34.3991 | -16.5374
9.6325 | -34.3995 | -16.5471
9.6386 | -34.3999 | -16.5567
9.6447 | -34.4003 | -16.5664
9.6508 | -34.4007 | -16.5759
9.6569 | -34.4011 | -16.5855
9.663 | -34.4015 | -16.5952
9.6691 | -34.402 | -16.6048
9.6753 | -34.4024 | -16.6144
9.6814 | -34.4028 | -16.6239
9.6875 | -34.4032 | -16.6334
9.6936 | -34.4036 | -16.6428
9.6997 | -34.4041 | -16.6526
9.7058 | -34.4045 | -16.6621
9.7119 | -34.4049 | -16.6719
9.718 | -34.4053 | -16.6816
9.7241 | -34.4058 | -16.6911
9.7302 | -34.4061 | -16.7008
9.7363 | -34.4066 | -16.7104
9.7424 | -34.407 | -16.7199
9.7485 | -34.4074 | -16.7296
9.7546 | -34.4078 | -16.7391
9.7607 | -34.4082 | -16.7489
9.7668 | -34.4087 | -16.7585
9.7729 | -34.4091 | -16.7682
9.779 | -34.4095 | -16.7776
9.7851 | -34.41 | -16.7873
9.7912 | -34.4104 | -16.7968
9.7973 | -34.4109 | -16.8066
9.8034 | -34.4113 | -16.816
9.8095 | -34.4117 | -16.8258
9.8157 | -34.4121 | -16.8355
9.8218 | -34.4125 | -16.8453
9.8279 | -34.4129 | -16.8548
9.834 | -34.4133 | -16.8645
9.8401 | -34.4138 | -16.874
9.8462 | -34.4142 | -16.8835
9.8523 | -34.4146 | -16.8931
9.8584 | -34.415 | -16.9027
9.8645 | -34.4155 | -16.9123
9.8706 | -34.4158 | -16.9219
9.8767 | -34.4162 | -16.9314
9.8828 | -34.4167 | -16.9409
9.8889 | -34.4171 | -16.9503
9.895 | -34.4175 | -16.9598
9.9011 | -34.4179 | -16.9694
9.9072 | -34.4183 | -16.979
9.9133 | -34.4187 | -16.9886
9.9194 | -34.4192 | -16.9981
9.9255 | -34.4196 | -17.0076
9.9316 | -34.42 | -17.017
9.9377 | -34.4204 | -17.0266
9.9438 | -34.4208 | -17.0363
9.9499 | -34.4213 | -17.0461
9.956 | -34.4217 | -17.0555
9.9622 | -34.4221 | -17.0651
9.9683 | -34.4225 | -17.0747
9.9744 | -34.4229 | -17.0844
9.9805 | -34.4234 | -17.094
9.9866 | -34.4237 | -17.1035
9.9927 | -34.4241 | -17.1129
9.9988 | -34.4245 | -17.1225
10.0049 | -34.425 | -17.132
10.011 | -34.4254 | -17.1417
10.0171 | -34.4258 | -17.1511
10.0232 | -34.4263 | -17.1606
10.0293 | -34.4267 | -17.1702
10.0354 | -34.4272 | -17.1798
10.0415 | -34.4276 | -17.1894
10.0476 | -34.428 | -17.199
10.0537 | -34.4284 | -17.2087
10.0598 | -34.4288 | -17.2183
10.0659 | -34.4292 | -17.228
10.072 | -34.4296 | -17.2373
10.0781 | -34.43 | -17.2468
10.0842 | -34.4305 | -17.2564
10.0903 | -34.4309 | -17.2659
10.0964 | -34.4313 | -17.2752
10.1026 | -34.4317 | -17.2848
10.1087 | -34.4321 | -17.2943
10.1148 | -34.4325 | -17.3039
10.1209 | -34.4329 | -17.3132
10.127 | -34.4333 | -17.3227
10.1331 | -34.4337 | -17.3321
10.1392 | -34.4341 | -17.3415
10.1453 | -34.4346 | -17.351
10.1514 | -34.435 | -17.3606
10.1575 | -34.4354 | -17.3703
10.1636 | -34.4358 | -17.3796
10.1697 | -34.4363 | -17.3891
10.1758 | -34.4367 | -17.3984
10.1819 | -34.4371 | -17.408
10.188 | -34.4375 | -17.4175
10.1941 | -34.4379 | -17.427
10.2002 | -34.4383 | -17.4366
10.2063 | -34.4387 | -17.446
10.2124 | -34.4391 | -17.4556
10.2185 | -34.4396 | -17.4652
10.2246 | -34.44 | -17.4745
10.2307 | -34.4404 | -17.4841
10.2368 | -34.4408 | -17.4934
10.2429 | -34.4412 | -17.5029
10.2491 | -34.4416 | -17.5125
10.2552 | -34.442 | -17.522
10.2613 | -34.4424 | -17.5316
10.2674 | -34.4428 | -17.5412
10.2735 | -34.4432 | -17.5507
10.2796 | -34.4436 | -17.5602
10.2857 | -34.4441 | -17.5698
10.2918 | -34.4444 | -17.5795
10.2979 | -34.4449 | -17.589
10.304 | -34.4453 | -17.5982
10.3101 | -34.4457 | -17.6078
10.3162 | -34.4461 | -17.6174
10.3223 | -34.4466 | -17.6268
10.3284 | -34.4469 | -17.6363
10.3345 | -34.4474 | -17.6457
10.3406 | -34.4478 | -17.6552
10.3467 | -34.4481 | -17.6646
10.3528 | -34.4486 | -17.674
10.3589 | -34.4489 | -17.6834
10.365 | -34.4494 | -17.6931
10.3711 | -34.4498 | -17.7025
10.3772 | -34.4502 | -17.712
10.3833 | -34.4506 | -17.7215
10.3895 | -34.451 | -17.7309
10.3956 | -34.4514 | -17.7404
10.4017 | -34.4518 | -17.7496
10.4078 | -34.4522 | -17.7592
10.4139 | -34.4526 | -17.7688
10.42 | -34.4531 | -17.7785
10.4261 | -34.4535 | -17.7881
10.4322 | -34.4539 | -17.7975
10.4383 | -34.4542 | -17.8068
10.4444 | -34.4546 | -17.8163
10.4505 | -34.455 | -17.8256
10.4566 | -34.4554 | -17.8349
10.4627 | -34.4559 | -17.8444
10.4688 | -34.4563 | -17.8537
10.4749 | -34.4567 | -17.8632
10.481 | -34.4571 | -17.8726
10.4871 | -34.4575 | -17.8821
10.4932 | -34.4579 | -17.8915
10.4993 | -34.4583 | -17.901
10.5054 | -34.4587 | -17.9104
10.5115 | -34.4591 | -17.9197
10.5176 | -34.4596 | -17.9289
10.5237 | -34.46 | -17.9385
10.5298 | -34.4604 | -17.9479
10.536 | -34.4608 | -17.9573
10.5421 | -34.4612 | -17.9668
10.5482 | -34.4615 | -17.9766
10.5543 | -34.4619 | -17.9858
10.5604 | -34.4623 | -17.9947
10.5665 | -34.4627 | -18.0045
10.5726 | -34.4631 | -18.0139
10.5787 | -34.4635 | -18.0233
10.5848 | -34.4639 | -18.0325
10.5909 | -34.4643 | -18.0421
10.597 | -34.4647 | -18.0515
10.6031 | -34.4651 | -18.0612
10.6092 | -34.4655 | -18.0705
10.6153 | -34.4659 | -18.0803
10.6214 | -34.4663 | -18.0897
10.6275 | -34.4667 | -18.0993
10.6336 | -34.4671 | -18.1087
10.6397 | -34.4675 | -18.1183
10.6458 | -34.468 | -18.1276
10.6519 | -34.4683 | -18.1372
10.658 | -34.4687 | -18.1467
10.6641 | -34.4691 | -18.1562
10.6702 | -34.4695 | -18.1658
10.6764 | -34.47 | -18.1752
10.6825 | -34.4704 | -18.1846
10.6886 | -34.4707 | -18.1936
10.6947 | -34.4711 | -18.2027
10.7008 | -34.4715 | -18.2123
10.7069 | -34.4719 | -18.2216
10.713 | -34.4722 | -18.2309
10.7191 | -34.4727 | -18.2401
10.7252 | -34.473 | -18.2491
10.7313 | -34.4734 | -18.2585
10.7374 | -34.4738 | -18.2679
10.7435 | -34.4741 | -18.2774
10.7496 | -34.4745 | -18.2865
10.7557 | -34.4749 | -18.2959
10.7618 | -34.4753 | -18.3046
10.7679 | -34.4757 | -18.3141
10.774 | -34.4761 | -18.3233
10.7801 | -34.4765 | -18.3329
10.7862 | -34.4769 | -18.342
10.7923 | -34.4773 | -18.3516
10.7984 | -34.4777 | -18.3604
10.8045 | -34.4781 | -18.3695
10.8106 | -34.4784 | -18.3789
10.8168 | -34.4788 | -18.3881
10.8229 | -34.4792 | -18.3976
10.829 | -34.4796 | -18.4068
10.8351 | -34.48 | -18.4157
10.8412 | -34.4804 | -18.4249
10.8473 | -34.4808 | -18.4338
10.8534 | -34.4812 | -18.4431
10.8595 | -34.4816 | -18.4525
10.8656 | -34.482 | -18.4617
10.8717 | -34.4824 | -18.4711
10.8778 | -34.4827 | -18.4806
10.8839 | -34.4831 | -18.4898
10.89 | -34.4835 | -18.4992
10.8961 | -34.4839 | -18.5085
10.9022 | -34.4843 | -18.5177
10.9083 | -34.4847 | -18.5268
10.9144 | -34.485 | -18.5361
10.9205 | -34.4854 | -18.5453
10.9266 | -34.4858 | -18.5545
10.9327 | -34.4862 | -18.5638
10.9388 | -34.4867 | -18.5728
10.9449 | -34.487 | -18.5819
10.951 | -34.4874 | -18.5912
10.9571 | -34.4878 | -18.6005
10.9633 | -34.4882 | -18.6096
10.9694 | -34.4886 | -18.619
10.9755 | -34.4889 | -18.6281
10.9816 | -34.4893 | -18.6375
10.9877 | -34.4898 | -18.6466
10.9938 | -34.4902 | -18.6561
10.9999 | -34.4906 | -18.6651
11.006 | -34.491 | -18.6742
11.0121 | -34.4914 | -18.6835
11.0182 | -34.4917 | -18.6925
11.0243 | -34.4921 | -18.7016
11.0304 | -34.4925 | -18.7109
11.0365 | -34.4929 | -18.7202
11.0426 | -34.4933 | -18.7293
11.0487 | -34.4936 | -18.7386
11.0548 | -34.494 | -18.7474
11.0609 | -34.4945 | -18.7566
11.067 | -34.4949 | -18.7657
11.0731 | -34.4953 | -18.7746
11.0792 | -34.4957 | -18.7836
11.0853 | -34.496 | -18.7929
11.0914 | -34.4965 | -18.8022
11.0975 | -34.4969 | -18.8115
11.1037 | -34.4972 | -18.8205
11.1098 | -34.4976 | -18.8297
11.1159 | -34.498 | -18.8389
11.122 | -34.4984 | -18.8481
11.1281 | -34.4988 | -18.8573
11.1342 | -34.4992 | -18.8665
11.1403 | -34.4996 | -18.8758
11.1464 | -34.5 | -18.8851
11.1525 | -34.5004 | -18.8942
11.1586 | -34.5008 | -18.9036
11.1647 | -34.5012 | -18.9127
11.1708 | -34.5016 | -18.9218
11.1769 | -34.5019 | -18.9309
11.183 | -34.5024 | -18.9395
11.1891 | -34.5028 | -18.9488
11.1952 | -34.5032 | -18.9577
11.2013 | -34.5036 | -18.9669
11.2074 | -34.504 | -18.9759
11.2135 | -34.5044 | -18.9851
11.2196 | -34.5049 | -18.994
11.2257 | -34.5053 | -19.0034
11.2318 | -34.5056 | -19.0125
11.2379 | -34.506 | -19.0215
11.244 | -34.5064 | -19.0311
11.2502 | -34.5069 | -19.0402
11.2563 | -34.5072 | -19.0493
11.2624 | -34.5076 | -19.0584
11.2685 | -34.508 | -19.0674
11.2746 | -34.5084 | -19.0764
11.2807 | -34.5088 | -19.0858
11.2868 | -34.5092 | -19.0951
11.2929 | -34.5096 | -19.1037
11.299 | -34.5101 | -19.1129
11.3051 | -34.5105 | -19.1222
11.3112 | -34.5109 | -19.1315
11.3173 | -34.5113 | -19.1404
11.3234 | -34.5117 | -19.1494
11.3295 | -34.5121 | -19.1586
11.3356 | -34.5124 | -19.1679
11.3417 | -34.5128 | -19.1768
11.3478 | -34.5132 | -19.1861
11.3539 | -34.5136 | -19.1952
11.36 | -34.514 | -19.2038
11.3661 | -34.5144 | -19.2127
11.3722 | -34.5148 | -19.2217
11.3783 | -34.5152 | -19.2308
11.3844 | -34.5157 | -19.24
11.3906 | -34.5161 | -19.2491
11.3967 | -34.5165 | -19.2579
11.4028 | -34.5169 | -19.267
11.4089 | -34.5173 | -19.2759
11.415 | -34.5177 | -19.285
11.4211 | -34.5181 | -19.2941
11.4272 | -34.5185 | -19.3031
11.4333 | -34.5188 | -19.3119
11.4394 | -34.5192 | -19.3207
11.4455 | -34.5195 | -19.3297
11.4516 | -34.5199 | -19.3387
11.4577 | -34.5203 | -19.3477
11.4638 | -34.5207 | -19.3564
11.4699 | -34.5212 | -19.3656
11.476 | -34.5216 | -19.3748
11.4821 | -34.522 | -19.3839
11.4882 | -34.5224 | -19.3931
11.4943 | -34.5228 | -19.4021
11.5004 | -34.5232 | -19.4109
11.5065 | -34.5236 | -19.42
11.5126 | -34.524 | -19.4288
11.5187 | -34.5244 | -19.4379
11.5248 | -34.5248 | -19.4472
11.5309 | -34.5252 | -19.4562
11.5371 | -34.5256 | -19.4652
11.5432 | -34.526 | -19.4744
11.5493 | -34.5264 | -19.4835
11.5554 | -34.5268 | -19.4924
11.5615 | -34.5272 | -19.5015
11.5676 | -34.5276 | -19.5105
11.5737 | -34.528 | -19.5191
11.5798 | -34.5284 | -19.528
11.5859 | -34.5288 | -19.5371
11.592 | -34.5292 | -19.5463
11.5981 | -34.5296 | -19.5552
11.6042 | -34.5301 | -19.5643
11.6103 | -34.5304 | -19.5732
11.6164 | -34.5309 | -19.5821
11.6225 | -34.5313 | -19.591
11.6286 | -34.5316 | -19.6
11.6347 | -34.532 | -19.609
11.6408 | -34.5325 | -19.6178
11.6469 | -34.5329 | -19.6267
11.653 | -34.5333 | -19.6355
11.6591 | -34.5336 | -19.6445
11.6652 | -34.5341 | -19.6535
11.6713 | -34.5345 | -19.6623
11.6775 | -34.5349 | -19.6712
11.6836 | -34.5353 | -19.6801
11.6897 | -34.5358 | -19.689
11.6958 | -34.5362 | -19.6981
11.7019 | -34.5366 | -19.7071
11.708 | -34.537 | -19.716
11.7141 | -34.5374 | -19.7246
11.7202 | -34.5378 | -19.7335
11.7263 | -34.5382 | -19.7425
11.7324 | -34.5386 | -19.7516
11.7385 | -34.539 | -19.7606
11.7446 | -34.5393 | -19.7694
11.7507 | -34.5398 | -19.7781
11.7568 | -34.5402 | -19.787
11.7629 | -34.5406 | -19.7961
11.769 | -34.541 | -19.8049
11.7751 | -34.5414 | -19.8137
11.7812 | -34.5419 | -19.8224
11.7873 | -34.5423 | -19.8313
11.7934 | -34.5427 | -19.8402
11.7995 | -34.5431 | -19.8491
11.8056 | -34.5435 | -19.858
11.8117 | -34.5439 | -19.867
11.8178 | -34.5443 | -19.8755
11.824 | -34.5448 | -19.8844
11.8301 | -34.5452 | -19.893
11.8362 | -34.5456 | -19.9018
11.8423 | -34.546 | -19.9107
11.8484 | -34.5464 | -19.9197
11.8545 | -34.5468 | -19.9285
11.8606 | -34.5472 | -19.9373
11.8667 | -34.5476 | -19.9462
11.8728 | -34.548 | -19.9548
11.8789 | -34.5484 | -19.9636
11.885 | -34.5489 | -19.9725
11.8911 | -34.5493 | -19.9814
11.8972 | -34.5497 | -19.9902
11.9033 | -34.5501 | -19.999
11.9094 | -34.5505 | -20.0079
11.9155 | -34.551 | -20.0166
11.9216 | -34.5514 | -20.0253
11.9277 | -34.5518 | -20.0339
11.9338 | -34.5523 | -20.0429
11.9399 | -34.5527 | -20.0517
11.946 | -34.5531 | -20.0604
11.9521 | -34.5535 | -20.0692
11.9582 | -34.5539 | -20.0778
11.9644 | -34.5543 | -20.0865
11.9705 | -34.5548 | -20.0953
11.9766 | -34.5552 | -20.1041
11.9827 | -34.5557 | -20.1129
11.9888 | -34.5561 | -20.1216
11.9949 | -34.5566 | -20.1301
12.001 | -34.557 | -20.139
12.0071 | -34.5574 | -20.1477
12.0132 | -34.5578 | -20.1564
12.0193 | -34.5582 | -20.1653
12.0254 | -34.5586 | -20.1741
12.0315 | -34.5591 | -20.1824
12.0376 | -34.5595 | -20.1912
12.0437 | -34.5599 | -20.2
12.0498 | -34.5603 | -20.2089
12.0559 | -34.5607 | -20.2176
12.062 | -34.5612 | -20.2262
12.0681 | -34.5616 | -20.2351
12.0742 | -34.562 | -20.2437
12.0803 | -34.5625 | -20.2526
12.0864 | -34.5629 | -20.2615
12.0925 | -34.5633 | -20.2702
12.0986 | -34.5638 | -20.2791
12.1047 | -34.5642 | -20.2879
12.1109 | -34.5647 | -20.2966
12.117 | -34.5651 | -20.3053
12.1231 | -34.5656 | -20.314
12.1292 | -34.566 | -20.3227
12.1353 | -34.5664 | -20.3315
12.1414 | -34.5669 | -20.3402
12.1475 | -34.5673 | -20.3489
12.1536 | -34.5677 | -20.3577
12.1597 | -34.5681 | -20.3664
12.1658 | -34.5686 | -20.3751
12.1719 | -34.569 | -20.3836
12.178 | -34.5694 | -20.3923
12.1841 | -34.5699 | -20.4011
12.1902 | -34.5703 | -20.4099
12.1963 | -34.5708 | -20.4187
12.2024 | -34.5712 | -20.427
12.2085 | -34.5717 | -20.4357
12.2146 | -34.5721 | -20.4444
12.2207 | -34.5726 | -20.4531
12.2268 | -34.573 | -20.4618
12.2329 | -34.5734 | -20.4703
12.239 | -34.5739 | -20.4786
12.2451 | -34.5744 | -20.4872
12.2513 | -34.5748 | -20.4958
12.2574 | -34.5752 | -20.5046
12.2635 | -34.5757 | -20.5134
12.2696 | -34.5761 | -20.5219
12.2757 | -34.5766 | -20.5305
12.2818 | -34.577 | -20.5391
12.2879 | -34.5775 | -20.5479
12.294 | -34.5779 | -20.5564
12.3001 | -34.5784 | -20.5651
12.3062 | -34.5788 | -20.5736
12.3123 | -34.5793 | -20.582
12.3184 | -34.5797 | -20.5905
12.3245 | -34.5802 | -20.5991
12.3306 | -34.5806 | -20.6076
12.3367 | -34.5811 | -20.6162
12.3428 | -34.5815 | -20.625
12.3489 | -34.582 | -20.6332
12.355 | -34.5824 | -20.642
12.3611 | -34.5829 | -20.6508
12.3672 | -34.5833 | -20.6593
12.3733 | -34.5838 | -20.668
12.3794 | -34.5842 | -20.6763
12.3855 | -34.5846 | -20.6849
12.3916 | -34.5851 | -20.6936
12.3978 | -34.5856 | -20.7023
12.4039 | -34.586 | -20.7109
12.41 | -34.5864 | -20.7195
12.4161 | -34.5869 | -20.7278
12.4222 | -34.5874 | -20.7363
12.4283 | -34.5878 | -20.7447
12.4344 | -34.5883 | -20.7534
12.4405 | -34.5887 | -20.7621
12.4466 | -34.5892 | -20.7705
12.4527 | -34.5897 | -20.7791
12.4588 | -34.5901 | -20.7878
12.4649 | -34.5906 | -20.7965
12.471 | -34.591 | -20.8052
12.4771 | -34.5915 | -20.8137
12.4832 | -34.592 | -20.8223
12.4893 | -34.5925 | -20.8308
12.4954 | -34.5929 | -20.8392
12.5015 | -34.5933 | -20.8478
12.5076 | -34.5938 | -20.8565
12.5137 | -34.5943 | -20.8652
12.5198 | -34.5948 | -20.8737
12.5259 | -34.5952 | -20.8825
12.532 | -34.5957 | -20.8909
12.5382 | -34.5962 | -20.8995
12.5443 | -34.5966 | -20.9081
12.5504 | -34.5971 | -20.9167
12.5565 | -34.5976 | -20.9252
12.5626 | -34.598 | -20.934
12.5687 | -34.5985 | -20.9426
12.5748 | -34.599 | -20.9512
12.5809 | -34.5994 | -20.9597
12.587 | -34.5999 | -20.9683
12.5931 | -34.6004 | -20.9768
12.5992 | -34.6009 | -20.9853
12.6053 | -34.6013 | -20.994
12.6114 | -34.6018 | -21.0025
12.6175 | -34.6023 | -21.0112
12.6236 | -34.6027 | -21.0199
12.6297 | -34.6032 | -21.0284
12.6358 | -34.6036 | -21.0369
12.6419 | -34.6041 | -21.0455
12.648 | -34.6045 | -21.0541
12.6541 | -34.605 | -21.0626
12.6602 | -34.6055 | -21.0709
12.6663 | -34.6059 | -21.0794
12.6724 | -34.6065 | -21.0879
12.6785 | -34.6069 | -21.0965
12.6847 | -34.6074 | -21.105
12.6908 | -34.6078 | -21.1135
12.6969 | -34.6083 | -21.122
12.703 | -34.6088 | -21.1306
12.7091 | -34.6092 | -21.1392
12.7152 | -34.6097 | -21.1477
12.7213 | -34.6102 | -21.1564
12.7274 | -34.6107 | -21.165
12.7335 | -34.6112 | -21.1734
12.7396 | -34.6116 | -21.1817
12.7457 | -34.6121 | -21.1904
12.7518 | -34.6125 | -21.1989
12.7579 | -34.613 | -21.2075
12.764 | -34.6134 | -21.2161
12.7701 | -34.6139 | -21.2248
12.7762 | -34.6144 | -21.2332
12.7823 | -34.6149 | -21.2418
12.7884 | -34.6154 | -21.2502
12.7945 | -34.6158 | -21.2588
12.8006 | -34.6163 | -21.2671
12.8067 | -34.6168 | -21.2755
12.8128 | -34.6173 | -21.2839
12.8189 | -34.6177 | -21.2925
12.8251 | -34.6182 | -21.3011
12.8312 | -34.6186 | -21.3096
12.8373 | -34.6191 | -21.3179
12.8434 | -34.6196 | -21.3264
12.8495 | -34.6201 | -21.335
12.8556 | -34.6205 | -21.3434
12.8617 | -34.621 | -21.352
12.8678 | -34.6215 | -21.3607
12.8739 | -34.622 | -21.3689
12.88 | -34.6224 | -21.3775
12.8861 | -34.6229 | -21.3858
12.8922 | -34.6233 | -21.3944
12.8983 | -34.6238 | -21.4028
12.9044 | -34.6243 | -21.4112
12.9105 | -34.6248 | -21.4195
12.9166 | -34.6252 | -21.4278
12.9227 | -34.6257 | -21.4363
12.9288 | -34.6261 | -21.4447
12.9349 | -34.6266 | -21.4532
12.941 | -34.6271 | -21.4616
12.9471 | -34.6276 | -21.4701
12.9532 | -34.6281 | -21.4786
12.9593 | -34.6286 | -21.4874
12.9654 | -34.629 | -21.4959
12.9716 | -34.6295 | -21.5045
12.9777 | -34.63 | -21.5131
12.9838 | -34.6305 | -21.5214
12.9899 | -34.631 | -21.5299
12.996 | -34.6314 | -21.5384
13.0021 | -34.6319 | -21.5469
13.0082 | -34.6324 | -21.5553
13.0143 | -34.6329 | -21.5636
13.0204 | -34.6334 | -21.572
13.0265 | -34.6339 | -21.5806
13.0326 | -34.6343 | -21.5891
13.0387 | -34.6349 | -21.5977
13.0448 | -34.6353 | -21.6062
13.0509 | -34.6358 | -21.6144
13.057 | -34.6363 | -21.6229
13.0631 | -34.6368 | -21.6316
13.0692 | -34.6373 | -21.6401
13.0753 | -34.6377 | -21.6486
13.0814 | -34.6382 | -21.6572
13.0875 | -34.6387 | -21.6654
13.0936 | -34.6392 | -21.6739
13.0997 | -34.6397 | -21.6826
13.1058 | -34.6401 | -21.6912
13.112 | -34.6406 | -21.6995
13.1181 | -34.6412 | -21.708
13.1242 | -34.6416 | -21.7165
13.1303 | -34.6421 | -21.7251
13.1364 | -34.6425 | -21.7337
13.1425 | -34.643 | -21.7418
13.1486 | -34.6434 | -21.7503
13.1547 | -34.6439 | -21.7588
13.1608 | -34.6444 | -21.7673
13.1669 | -34.6449 | -21.7753
13.173 | -34.6454 | -21.7838
13.1791 | -34.6459 | -21.7922
13.1852 | -34.6463 | -21.8006
13.1913 | -34.6468 | -21.8087
13.1974 | -34.6473 | -21.8174
13.2035 | -34.6478 | -21.8258
13.2096 | -34.6483 | -21.8344
13.2157 | -34.6488 | -21.8427
13.2218 | -34.6493 | -21.851
13.2279 | -34.6498 | -21.8593
13.234 | -34.6503 | -21.8677
13.2401 | -34.6508 | -21.8763
13.2462 | -34.6513 | -21.8847
13.2524 | -34.6518 | -21.893
13.2585 | -34.6523 | -21.9011
13.2646 | -34.6527 | -21.9096
13.2707 | -34.6532 | -21.9182
13.2768 | -34.6537 | -21.9268
13.2829 | -34.6543 | -21.9354
13.289 | -34.6548 | -21.9438
13.2951 | -34.6553 | -21.9521
13.3012 | -34.6558 | -21.9604
13.3073 | -34.6563 | -21.9688
13.3134 | -34.6567 | -21.9771
13.3195 | -34.6572 | -21.9852
13.3256 | -34.6577 | -21.9934
13.3317 | -34.6583 | -22.0018
13.3378 | -34.6588 | -22.0104
13.3439 | -34.6592 | -22.0192
13.35 | -34.6597 | -22.028
13.3561 | -34.6602 | -22.0362
13.3622 | -34.6607 | -22.0447
13.3683 | -34.6612 | -22.0532
13.3744 | -34.6616 | -22.0615
13.3805 | -34.6621 | -22.0698
13.3866 | -34.6626 | -22.0778
13.3927 | -34.663 | -22.0862
13.3989 | -34.6635 | -22.0944
13.405 | -34.664 | -22.1031
13.4111 | -34.6645 | -22.1113
13.4172 | -34.6649 | -22.1197
13.4233 | -34.6654 | -22.1284
13.4294 | -34.6658 | -22.1372
13.4355 | -34.6663 | -22.1454
13.4416 | -34.6668 | -22.1537
13.4477 | -34.6673 | -22.1619
13.4538 | -34.6678 | -22.1703
13.4599 | -34.6683 | -22.1787
13.466 | -34.6687 | -22.1872
13.4721 | -34.6692 | -22.1957
13.4782 | -34.6697 | -22.204
13.4843 | -34.6702 | -22.2123
13.4904 | -34.6707 | -22.2206
13.4965 | -34.6713 | -22.2291
13.5026 | -34.6718 | -22.2375
13.5087 | -34.6723 | -22.2458
13.5148 | -34.6728 | -22.2545
13.5209 | -34.6733 | -22.2631
13.527 | -34.6738 | -22.2716
13.5331 | -34.6744 | -22.2801
13.5393 | -34.6748 | -22.288
13.5454 | -34.6753 | -22.2962
13.5515 | -34.6758 | -22.3049
13.5576 | -34.6763 | -22.3137
13.5637 | -34.6768 | -22.3223
13.5698 | -34.6772 | -22.3306
13.5759 | -34.6777 | -22.3395
13.582 | -34.6782 | -22.3482
13.5881 | -34.6787 | -22.3566
13.5942 | -34.6793 | -22.3651
13.6003 | -34.6797 | -22.3736
13.6064 | -34.6803 | -22.3823
13.6125 | -34.6807 | -22.3907
13.6186 | -34.6812 | -22.3989
13.6247 | -34.6817 | -22.4074
13.6308 | -34.6822 | -22.4158
13.6369 | -34.6828 | -22.4241
13.643 | -34.6832 | -22.4326
13.6491 | -34.6837 | -22.4407
13.6552 | -34.6843 | -22.4489
13.6613 | -34.6848 | -22.457
13.6674 | -34.6853 | -22.4651
13.6735 | -34.6857 | -22.4732
13.6796 | -34.6862 | -22.4815
13.6858 | -34.6867 | -22.4893
13.6919 | -34.6872 | -22.4974
13.698 | -34.6877 | -22.5059
13.7041 | -34.6882 | -22.5143
13.7102 | -34.6888 | -22.5226
13.7163 | -34.6893 | -22.531
13.7224 | -34.6898 | -22.5393
13.7285 | -34.6903 | -22.5476
13.7346 | -34.6908 | -22.5562
13.7407 | -34.6913 | -22.5645
13.7468 | -34.6918 | -22.5729
13.7529 | -34.6923 | -22.5812
13.759 | -34.6927 | -22.5897
13.7651 | -34.6932 | -22.5981
13.7712 | -34.6937 | -22.6065
13.7773 | -34.6942 | -22.615
13.7834 | -34.6947 | -22.6234
13.7895 | -34.6952 | -22.6316
13.7956 | -34.6957 | -22.6401
13.8017 | -34.6962 | -22.6483
13.8078 | -34.6967 | -22.6565
13.8139 | -34.6972 | -22.665
13.82 | -34.6977 | -22.6733
13.8262 | -34.6982 | -22.6816
13.8323 | -34.6987 | -22.69
13.8384 | -34.6992 | -22.6983
13.8445 | -34.6997 | -22.7069
13.8506 | -34.7002 | -22.7152
13.8567 | -34.7008 | -22.7239
13.8628 | -34.7013 | -22.7325
13.8689 | -34.7017 | -22.7409
13.875 | -34.7022 | -22.7494
13.8811 | -34.7027 | -22.7578
13.8872 | -34.7032 | -22.766
13.8933 | -34.7037 | -22.7741
13.8994 | -34.7042 | -22.7825
13.9055 | -34.7046 | -22.7909
13.9116 | -34.7052 | -22.7992
13.9177 | -34.7057 | -22.8074
13.9238 | -34.7061 | -22.8158
13.9299 | -34.7066 | -22.8241
13.936 | -34.7071 | -22.8326
13.9421 | -34.7076 | -22.841
13.9482 | -34.7081 | -22.8493
13.9543 | -34.7085 | -22.8576
13.9604 | -34.7091 | -22.866
13.9665 | -34.7096 | -22.8742
13.9727 | -34.7101 | -22.8826
13.9788 | -34.7106 | -22.8911
13.9849 | -34.7111 | -22.8995
13.991 | -34.7116 | -22.908
13.9971 | -34.7121 | -22.9164
14.0032 | -34.7126 | -22.9249
14.0093 | -34.7131 | -22.9333
14.0154 | -34.7136 | -22.9417
14.0215 | -34.7141 | -22.95
14.0276 | -34.7146 | -22.9582
14.0337 | -34.7151 | -22.9664
14.0398 | -34.7156 | -22.9749
14.0459 | -34.7161 | -22.9833
14.052 | -34.7166 | -22.9916
14.0581 | -34.7171 | -23
14.0642 | -34.7176 | -23.0085
14.0703 | -34.7181 | -23.0171
14.0764 | -34.7186 | -23.0255
14.0825 | -34.7191 | -23.0339
14.0886 | -34.7196 | -23.0422
14.0947 | -34.7201 | -23.0508
14.1008 | -34.7206 | -23.0591
14.1069 | -34.7211 | -23.0672
14.1131 | -34.7216 | -23.0758
14.1192 | -34.7221 | -23.0843
14.1253 | -34.7226 | -23.0924
14.1314 | -34.7231 | -23.1007
14.1375 | -34.7236 | -23.1089
14.1436 | -34.7241 | -23.1173
14.1497 | -34.7246 | -23.1258
14.1558 | -34.7251 | -23.1342
14.1619 | -34.7256 | -23.1425
14.168 | -34.7261 | -23.151
14.1741 | -34.7266 | -23.1593
14.1802 | -34.7271 | -23.1675
14.1863 | -34.7277 | -23.1758
14.1924 | -34.7282 | -23.1841
14.1985 | -34.7287 | -23.1926
14.2046 | -34.7292 | -23.201
14.2107 | -34.7297 | -23.2094
14.2168 | -34.7302 | -23.218
14.2229 | -34.7307 | -23.2264
14.229 | -34.7312 | -23.2347
14.2351 | -34.7317 | -23.243
14.2412 | -34.7322 | -23.2514
14.2473 | -34.7326 | -23.2596
14.2534 | -34.7331 | -23.2678
14.2596 | -34.7336 | -23.2759
14.2657 | -34.7341 | -23.2842
14.2718 | -34.7346 | -23.2925
14.2779 | -34.7351 | -23.3009
14.284 | -34.7356 | -23.3091
14.2901 | -34.7361 | -23.3175
14.2962 | -34.7367 | -23.3259
14.3023 | -34.7371 | -23.3341
14.3084 | -34.7376 | -23.3423
14.3145 | -34.7382 | -23.3506
14.3206 | -34.7386 | -23.3593
14.3267 | -34.7391 | -23.3677
14.3328 | -34.7396 | -23.376
14.3389 | -34.7401 | -23.3843
14.345 | -34.7406 | -23.3925
14.3511 | -34.7411 | -23.4007
14.3572 | -34.7416 | -23.4092
14.3633 | -34.7421 | -23.4176
14.3694 | -34.7426 | -23.426
14.3755 | -34.7431 | -23.4342
14.3816 | -34.7437 | -23.4428
14.3877 | -34.7442 | -23.4512
14.3938 | -34.7447 | -23.4596
14.4 | -34.7452 | -23.4678
14.4061 | -34.7457 | -23.4761
14.4122 | -34.7462 | -23.4845
14.4183 | -34.7467 | -23.4927
14.4244 | -34.7472 | -23.5012
14.4305 | -34.7478 | -23.5095
14.4366 | -34.7483 | -23.5177
14.4427 | -34.7488 | -23.5263
14.4488 | -34.7493 | -23.5347
14.4549 | -34.7498 | -23.5429
14.461 | -34.7503 | -23.5511
14.4671 | -34.7507 | -23.5596
14.4732 | -34.7513 | -23.5679
14.4793 | -34.7518 | -23.5763
14.4854 | -34.7523 | -23.5845
14.4915 | -34.7528 | -23.5925
14.4976 | -34.7533 | -23.6007
14.5037 | -34.7538 | -23.6088
14.5098 | -34.7543 | -23.6171
14.5159 | -34.7548 | -23.6255
14.522 | -34.7553 | -23.6337
14.5281 | -34.7557 | -23.6419
14.5342 | -34.7562 | -23.6503
14.5403 | -34.7567 | -23.6586
14.5465 | -34.7572 | -23.667
14.5526 | -34.7577 | -23.6756
14.5587 | -34.7582 | -23.6842
14.5648 | -34.7588 | -23.6924
14.5709 | -34.7592 | -23.7004
14.577 | -34.7597 | -23.7088
14.5831 | -34.7602 | -23.717
14.5892 | -34.7608 | -23.7252
14.5953 | -34.7613 | -23.7332
14.6014 | -34.7618 | -23.7416
14.6075 | -34.7623 | -23.75
14.6136 | -34.7628 | -23.7584
14.6197 | -34.7633 | -23.7667
14.6258 | -34.7638 | -23.7749
14.6319 | -34.7643 | -23.7831
14.638 | -34.7648 | -23.7913
14.6441 | -34.7652 | -23.7996
14.6502 | -34.7657 | -23.808
14.6563 | -34.7662 | -23.8164
14.6624 | -34.7667 | -23.8247
14.6685 | -34.7672 | -23.8326
14.6746 | -34.7677 | -23.8409
14.6807 | -34.7682 | -23.849
14.6869 | -34.7687 | -23.8573
14.693 | -34.7692 | -23.8656
14.6991 | -34.7697 | -23.8739
14.7052 | -34.7703 | -23.8821
14.7113 | -34.7708 | -23.8905
14.7174 | -34.7713 | -23.8988
14.7235 | -34.7718 | -23.907
14.7296 | -34.7723 | -23.9152
14.7357 | -34.7728 | -23.9234
14.7418 | -34.7733 | -23.9318
14.7479 | -34.7738 | -23.9401
14.754 | -34.7743 | -23.9485
14.7601 | -34.7748 | -23.9568
14.7662 | -34.7753 | -23.9652
14.7723 | -34.7758 | -23.9735
14.7784 | -34.7763 | -23.9819
14.7845 | -34.7768 | -23.9902
14.7906 | -34.7773 | -23.9985
14.7967 | -34.7778 | -24.0065
14.8028 | -34.7783 | -24.0147
14.8089 | -34.7788 | -24.023
14.815 | -34.7793 | -24.0313
14.8211 | -34.7798 | -24.0394
14.8272 | -34.7803 | -24.0478
14.8334 | -34.7808 | -24.0562
14.8395 | -34.7813 | -24.0643
14.8456 | -34.7818 | -24.0726
14.8517 | -34.7823 | -24.0809
14.8578 | -34.7828 | -24.0893
14.8639 | -34.7833 | -24.0977
14.87 | -34.7838 | -24.1059
14.8761 | -34.7843 | -24.1139
14.8822 | -34.7848 | -24.1223
14.8883 | -34.7852 | -24.1307
14.8944 | -34.7857 | -24.1387
14.9005 | -34.7862 | -24.147
14.9066 | -34.7867 | -24.155
14.9127 | -34.7872 | -24.1634
14.9188 | -34.7877 | -24.1716
14.9249 | -34.7882 | -24.1799
14.931 | -34.7888 | -24.1884
14.9371 | -34.7893 | -24.1968
14.9432 | -34.7898 | -24.205
14.9493 | -34.7903 | -24.2132
14.9554 | -34.7908 | -24.2216
14.9615 | -34.7913 | -24.2299
14.9676 | -34.7917 | -24.2383
14.9738 | -34.7922 | -24.2464
14.9799 | -34.7928 | -24.2546
14.986 | -34.7932 | -24.2629
14.9921 | -34.7937 | -24.2709
14.9982 | -34.7942 | -24.2795
15.0043 | -34.7947 | -24.2877
15.0104 | -34.7952 | -24.2962
15.0165 | -34.7957 | -24.3044
15.0226 | -34.7962 | -24.3125
15.0287 | -34.7967 | -24.321
15.0348 | -34.7972 | -24.3292
15.0409 | -34.7976 | -24.3373
15.047 | -34.7981 | -24.3456
15.0531 | -34.7987 | -24.3538
15.0592 | -34.7991 | -24.362
15.0653 | -34.7996 | -24.3702
15.0714 | -34.8001 | -24.3783
15.0775 | -34.8006 | -24.3864
15.0836 | -34.8011 | -24.3943
15.0897 | -34.8016 | -24.4024
15.0958 | -34.8021 | -24.4105
15.1019 | -34.8026 | -24.4188
15.108 | -34.8031 | -24.4269
15.1141 | -34.8036 | -24.4352
15.1203 | -34.8041 | -24.4434
15.1264 | -34.8046 | -24.4515
15.1325 | -34.8051 | -24.4596
15.1386 | -34.8056 | -24.4677
15.1447 | -34.8061 | -24.476
15.1508 | -34.8066 | -24.4841
15.1569 | -34.8071 | -24.4923
15.163 | -34.8075 | -24.5003
15.1691 | -34.808 | -24.5083
15.1752 | -34.8085 | -24.5165
15.1813 | -34.809 | -24.5247
15.1874 | -34.8095 | -24.5332
15.1935 | -34.81 | -24.5413
15.1996 | -34.8105 | -24.5494
15.2057 | -34.811 | -24.5578
15.2118 | -34.8115 | -24.566
15.2179 | -34.8121 | -24.5744
15.224 | -34.8125 | -24.5826
15.2301 | -34.813 | -24.5907
15.2362 | -34.8135 | -24.5992
15.2423 | -34.814 | -24.6078
15.2484 | -34.8145 | -24.6159
15.2545 | -34.815 | -24.624
15.2607 | -34.8155 | -24.6322
15.2668 | -34.816 | -24.6401
15.2729 | -34.8165 | -24.6484
15.279 | -34.817 | -24.6567
15.2851 | -34.8175 | -24.6649
15.2912 | -34.818 | -24.6732
15.2973 | -34.8185 | -24.6816
15.3034 | -34.819 | -24.6897
15.3095 | -34.8194 | -24.6979
15.3156 | -34.82 | -24.706
15.3217 | -34.8204 | -24.7138
15.3278 | -34.8209 | -24.7222
15.3339 | -34.8214 | -24.7301
15.34 | -34.8219 | -24.7381
15.3461 | -34.8224 | -24.7463
15.3522 | -34.8229 | -24.7546
15.3583 | -34.8233 | -24.7629
15.3644 | -34.8238 | -24.771
15.3705 | -34.8243 | -24.7788
15.3766 | -34.8248 | -24.7871
15.3827 | -34.8253 | -24.7951
15.3888 | -34.8258 | -24.8032
15.3949 | -34.8263 | -24.8114
15.401 | -34.8268 | -24.8196
15.4072 | -34.8273 | -24.8276
15.4133 | -34.8277 | -24.8357
15.4194 | -34.8282 | -24.844
15.4255 | -34.8287 | -24.8521
15.4316 | -34.8292 | -24.8602
15.4377 | -34.8296 | -24.8684
15.4438 | -34.8301 | -24.8765
15.4499 | -34.8306 | -24.8845
15.456 | -34.8312 | -24.8924
15.4621 | -34.8317 | -24.9004
15.4682 | -34.8322 | -24.9088
15.4743 | -34.8326 | -24.9167
15.4804 | -34.8331 | -24.9249
15.4865 | -34.8335 | -24.9331
15.4926 | -34.834 | -24.9414
15.4987 | -34.8345 | -24.9495
15.5048 | -34.835 | -24.9576
15.5109 | -34.8355 | -24.9658
15.517 | -34.836 | -24.9741
15.5231 | -34.8365 | -24.9823
15.5292 | -34.8369 | -24.9904
15.5353 | -34.8374 | -24.9987
15.5414 | -34.838 | -25.0068
15.5476 | -34.8384 | -25.0146
15.5537 | -34.8389 | -25.0225
15.5598 | -34.8394 | -25.031
15.5659 | -34.8399 | -25.039
15.572 | -34.8404 | -25.0469
15.5781 | -34.8408 | -25.0552
15.5842 | -34.8413 | -25.0632
15.5903 | -34.8418 | -25.0712
15.5964 | -34.8422 | -25.0793
15.6025 | -34.8427 | -25.0873
15.6086 | -34.8432 | -25.0953
15.6147 | -34.8437 | -25.1034
15.6208 | -34.8441 | -25.1112
15.6269 | -34.8446 | -25.1192
15.633 | -34.8451 | -25.1272
15.6391 | -34.8456 | -25.1351
15.6452 | -34.8461 | -25.1434
15.6513 | -34.8465 | -25.1513
15.6574 | -34.847 | -25.1594
15.6635 | -34.8475 | -25.1677
15.6696 | -34.848 | -25.1758
15.6757 | -34.8485 | -25.184
15.6818 | -34.849 | -25.1919
15.688 | -34.8494 | -25.1999
15.6941 | -34.8499 | -25.2078
15.7002 | -34.8504 | -25.2159
15.7063 | -34.8508 | -25.224
15.7124 | -34.8513 | -25.2321
15.7185 | -34.8518 | -25.2403
15.7246 | -34.8522 | -25.2482
15.7307 | -34.8527 | -25.2562
15.7368 | -34.8532 | -25.2646
15.7429 | -34.8537 | -25.2726
15.749 | -34.8542 | -25.2806
15.7551 | -34.8546 | -25.2887
15.7612 | -34.8551 | -25.2966
15.7673 | -34.8556 | -25.3046
15.7734 | -34.856 | -25.3126
15.7795 | -34.8565 | -25.3207
15.7856 | -34.8569 | -25.3287
15.7917 | -34.8574 | -25.3363
15.7978 | -34.8579 | -25.3443
15.8039 | -34.8584 | -25.3523
15.81 | -34.8588 | -25.3602
15.8161 | -34.8593 | -25.3682
15.8222 | -34.8597 | -25.3762
15.8283 | -34.8602 | -25.384
15.8345 | -34.8607 | -25.3919
15.8406 | -34.8612 | -25.3999
15.8467 | -34.8617 | -25.4078
15.8528 | -34.8621 | -25.4159
15.8589 | -34.8626 | -25.4239
15.865 | -34.8631 | -25.4318
15.8711 | -34.8636 | -25.4398
15.8772 | -34.864 | -25.4476
15.8833 | -34.8645 | -25.4558
15.8894 | -34.8649 | -25.464
15.8955 | -34.8654 | -25.4721
15.9016 | -34.8658 | -25.48
15.9077 | -34.8663 | -25.488
15.9138 | -34.8668 | -25.4962
15.9199 | -34.8672 | -25.504
15.926 | -34.8677 | -25.5119
15.9321 | -34.8682 | -25.5197
15.9382 | -34.8686 | -25.5276
15.9443 | -34.8691 | -25.5356
15.9504 | -34.8695 | -25.5436
15.9565 | -34.87 | -25.5512
15.9626 | -34.8705 | -25.5593
15.9687 | -34.8709 | -25.5668
15.9749 | -34.8714 | -25.5748
15.981 | -34.8719 | -25.5825
15.9871 | -34.8724 | -25.5906
15.9932 | -34.8728 | -25.5986
15.9993 | -34.8733 | -25.6066
16.0054 | -34.8738 | -25.6143
16.0115 | -34.8743 | -25.6224
16.0176 | -34.8748 | -25.6304
16.0237 | -34.8752 | -25.6385
16.0298 | -34.8757 | -25.6461
16.0359 | -34.8761 | -25.6539
16.042 | -34.8766 | -25.6617
16.0481 | -34.8771 | -25.6696
16.0542 | -34.8776 | -25.6777
16.0603 | -34.878 | -25.6855
16.0664 | -34.8785 | -25.6935
16.0725 | -34.8789 | -25.701
16.0786 | -34.8794 | -25.7087
16.0847 | -34.8798 | -25.7168
16.0908 | -34.8803 | -25.7248
16.0969 | -34.8808 | -25.7327
16.103 | -34.8812 | -25.7406
16.1091 | -34.8817 | -25.7485
16.1152 | -34.8822 | -25.7563
16.1214 | -34.8826 | -25.7643
16.1275 | -34.8831 | -25.7722
16.1336 | -34.8836 | -25.7801
16.1397 | -34.8841 | -25.7881
16.1458 | -34.8845 | -25.796
16.1519 | -34.885 | -25.8041
16.158 | -34.8854 | -25.8121
16.1641 | -34.8859 | -25.8201
16.1702 | -34.8863 | -25.8281
16.1763 | -34.8868 | -25.8359
16.1824 | -34.8873 | -25.8437
16.1885 | -34.8877 | -25.8515
16.1946 | -34.8882 | -25.8592
16.2007 | -34.8887 | -25.8674
16.2068 | -34.8892 | -25.8752
16.2129 | -34.8896 | -25.8832
16.219 | -34.8901 | -25.8908
16.2251 | -34.8905 | -25.8987
16.2312 | -34.891 | -25.9064
16.2373 | -34.8915 | -25.9141
16.2434 | -34.892 | -25.9222
16.2495 | -34.8925 | -25.9299
16.2556 | -34.893 | -25.9378
16.2618 | -34.8935 | -25.9457
16.2679 | -34.894 | -25.9537
16.274 | -34.8945 | -25.9615
16.2801 | -34.895 | -25.9693
16.2862 | -34.8954 | -25.9771
16.2923 | -34.8959 | -25.9849
16.2984 | -34.8963 | -25.9929
16.3045 | -34.8968 | -26.0007
16.3106 | -34.8972 | -26.0086
16.3167 | -34.8977 | -26.0163
16.3228 | -34.8982 | -26.0239
16.3289 | -34.8986 | -26.0318
16.335 | -34.8991 | -26.0396
16.3411 | -34.8995 | -26.0474
16.3472 | -34.9 | -26.0554
16.3533 | -34.9005 | -26.0632
16.3594 | -34.9009 | -26.0709
16.3655 | -34.9014 | -26.0787
16.3716 | -34.9018 | -26.0867
16.3777 | -34.9023 | -26.0942
16.3838 | -34.9028 | -26.102
16.3899 | -34.9033 | -26.1094
16.396 | -34.9037 | -26.1173
16.4021 | -34.9042 | -26.125
16.4083 | -34.9047 | -26.1327
16.4144 | -34.9051 | -26.1405
16.4205 | -34.9056 | -26.1484
16.4266 | -34.906 | -26.1561
16.4327 | -34.9065 | -26.1639
16.4388 | -34.9069 | -26.172
16.4449 | -34.9074 | -26.1797
16.451 | -34.9079 | -26.1875
16.4571 | -34.9084 | -26.1954
16.4632 | -34.9088 | -26.203
16.4693 | -34.9092 | -26.2108
16.4754 | -34.9097 | -26.2185
16.4815 | -34.9101 | -26.2264
16.4876 | -34.9106 | -26.2337
16.4937 | -34.9111 | -26.2413
16.4998 | -34.9116 | -26.2487
16.5059 | -34.9121 | -26.2563
16.512 | -34.9125 | -26.2643
16.5181 | -34.913 | -26.2721
16.5242 | -34.9135 | -26.28
16.5303 | -34.914 | -26.2876
16.5364 | -34.9144 | -26.2952
16.5425 | -34.9149 | -26.3029
16.5487 | -34.9153 | -26.3106
16.5548 | -34.9158 | -26.3186
16.5609 | -34.9163 | -26.3266
16.567 | -34.9168 | -26.3342
16.5731 | -34.9173 | -26.3422
16.5792 | -34.9177 | -26.3497
16.5853 | -34.9182 | -26.3575
16.5914 | -34.9187 | -26.3654
16.5975 | -34.9191 | -26.3733
16.6036 | -34.9196 | -26.3809
16.6097 | -34.92 | -26.3889
16.6158 | -34.9204 | -26.3966
16.6219 | -34.9209 | -26.4042
16.628 | -34.9213 | -26.4121
16.6341 | -34.9218 | -26.4199
16.6402 | -34.9222 | -26.4276
16.6463 | -34.9227 | -26.4355
16.6524 | -34.9232 | -26.4431
16.6585 | -34.9237 | -26.4511
16.6646 | -34.9241 | -26.4588
16.6707 | -34.9246 | -26.4665
16.6768 | -34.9251 | -26.4743
16.6829 | -34.9256 | -26.4821
16.689 | -34.9261 | -26.4898
16.6952 | -34.9265 | -26.4976
16.7013 | -34.927 | -26.5053
16.7074 | -34.9274 | -26.5128
16.7135 | -34.9279 | -26.5205
16.7196 | -34.9283 | -26.5282
16.7257 | -34.9288 | -26.536
16.7318 | -34.9293 | -26.5437
16.7379 | -34.9298 | -26.5513
16.744 | -34.9302 | -26.5587
16.7501 | -34.9307 | -26.5666
16.7562 | -34.9311 | -26.5745
16.7623 | -34.9315 | -26.582
16.7684 | -34.932 | -26.5897
16.7745 | -34.9325 | -26.5975
16.7806 | -34.933 | -26.6052
16.7867 | -34.9335 | -26.6129
16.7928 | -34.9339 | -26.6208
16.7989 | -34.9343 | -26.6286
16.805 | -34.9348 | -26.6365
16.8111 | -34.9353 | -26.6439
16.8172 | -34.9358 | -26.652
16.8233 | -34.9363 | -26.6597
16.8294 | -34.9368 | -26.6674
16.8356 | -34.9373 | -26.6751
16.8417 | -34.9378 | -26.6826
16.8478 | -34.9383 | -26.6902
16.8539 | -34.9387 | -26.698
16.86 | -34.9392 | -26.7057
16.8661 | -34.9397 | -26.7135
16.8722 | -34.9401 | -26.7214
16.8783 | -34.9406 | -26.7292
16.8844 | -34.9411 | -26.7367
16.8905 | -34.9415 | -26.7445
16.8966 | -34.942 | -26.752
16.9027 | -34.9425 | -26.7594
16.9088 | -34.943 | -26.7672
16.9149 | -34.9434 | -26.7749
16.921 | -34.944 | -26.7822
16.9271 | -34.9444 | -26.79
16.9332 | -34.9449 | -26.7976
16.9393 | -34.9454 | -26.8053
16.9454 | -34.9459 | -26.8129
16.9515 | -34.9463 | -26.8208
16.9576 | -34.9468 | -26.8284
16.9637 | -34.9473 | -26.8359
16.9698 | -34.9478 | -26.8434
16.9759 | -34.9482 | -26.851
16.9821 | -34.9487 | -26.8589
16.9882 | -34.9492 | -26.8663
16.9943 | -34.9497 | -26.8739
17.0004 | -34.9502 | -26.8816
17.0065 | -34.9507 | -26.8892
17.0126 | -34.9511 | -26.8968
17.0187 | -34.9516 | -26.9045
17.0248 | -34.9521 | -26.9119
17.0309 | -34.9525 | -26.9195
17.037 | -34.953 | -26.9271
17.0431 | -34.9535 | -26.9347
17.0492 | -34.954 | -26.9424
17.0553 | -34.9544 | -26.9501
17.0614 | -34.955 | -26.9576
17.0675 | -34.9554 | -26.9653
17.0736 | -34.9559 | -26.973
17.0797 | -34.9563 | -26.9805
17.0858 | -34.9568 | -26.9882
17.0919 | -34.9573 | -26.9959
17.098 | -34.9577 | -27.0034
17.1041 | -34.9582 | -27.0112
17.1102 | -34.9586 | -27.0187
17.1163 | -34.9591 | -27.0264
17.1225 | -34.9596 | -27.0337
17.1286 | -34.9601 | -27.0412
17.1347 | -34.9606 | -27.0486
17.1408 | -34.9611 | -27.0563
17.1469 | -34.9615 | -27.0639
17.153 | -34.962 | -27.0714
17.1591 | -34.9624 | -27.0789
17.1652 | -34.963 | -27.0864
17.1713 | -34.9634 | -27.0941
17.1774 | -34.9639 | -27.1016
17.1835 | -34.9644 | -27.1093
17.1896 | -34.9648 | -27.1167
17.1957 | -34.9653 | -27.1241
17.2018 | -34.9658 | -27.1315
17.2079 | -34.9663 | -27.1391
17.214 | -34.9668 | -27.1467
17.2201 | -34.9672 | -27.1541
17.2262 | -34.9677 | -27.1619
17.2323 | -34.9682 | -27.1696
17.2384 | -34.9687 | -27.1771
17.2445 | -34.9692 | -27.1847
17.2506 | -34.9696 | -27.1921
17.2567 | -34.9701 | -27.1998
17.2628 | -34.9706 | -27.2073
17.269 | -34.9711 | -27.2147
17.2751 | -34.9715 | -27.2222
17.2812 | -34.972 | -27.23
17.2873 | -34.9725 | -27.2375
17.2934 | -34.973 | -27.2452
17.2995 | -34.9735 | -27.2528
17.3056 | -34.974 | -27.26
17.3117 | -34.9745 | -27.2679
17.3178 | -34.9749 | -27.2755
17.3239 | -34.9754 | -27.2831
17.33 | -34.9759 | -27.2908
17.3361 | -34.9763 | -27.2983
17.3422 | -34.9769 | -27.3058
17.3483 | -34.9773 | -27.3134
17.3544 | -34.9778 | -27.3212
17.3605 | -34.9783 | -27.3285
17.3666 | -34.9787 | -27.3361
17.3727 | -34.9792 | -27.3437
17.3788 | -34.9797 | -27.3511
17.3849 | -34.9802 | -27.3586
17.391 | -34.9807 | -27.3659
17.3971 | -34.9811 | -27.3734
17.4032 | -34.9816 | -27.3809
17.4094 | -34.9821 | -27.3884
17.4155 | -34.9825 | -27.396
17.4216 | -34.983 | -27.4033
17.4277 | -34.9835 | -27.4108
17.4338 | -34.984 | -27.4184
17.4399 | -34.9845 | -27.4262
17.446 | -34.985 | -27.4336
17.4521 | -34.9855 | -27.441
17.4582 | -34.986 | -27.4485
17.4643 | -34.9865 | -27.4558
17.4704 | -34.9869 | -27.4633
17.4765 | -34.9874 | -27.4711
17.4826 | -34.9879 | -27.4781
17.4887 | -34.9884 | -27.4858
17.4948 | -34.9889 | -27.4933
17.5009 | -34.9893 | -27.5008
17.507 | -34.9898 | -27.5084
17.5131 | -34.9903 | -27.5159
17.5192 | -34.9908 | -27.523
17.5253 | -34.9913 | -27.5305
17.5314 | -34.9917 | -27.538
17.5375 | -34.9922 | -27.5455
17.5436 | -34.9927 | -27.5532
17.5497 | -34.9932 | -27.5607
17.5559 | -34.9937 | -27.5678
17.562 | -34.9941 | -27.5754
17.5681 | -34.9946 | -27.5829
17.5742 | -34.9951 | -27.5906
17.5803 | -34.9956 | -27.5981
17.5864 | -34.9961 | -27.6055
17.5925 | -34.9966 | -27.6129
17.5986 | -34.9971 | -27.6204
17.6047 | -34.9975 | -27.628
17.6108 | -34.998 | -27.6354
17.6169 | -34.9985 | -27.6429
17.623 | -34.999 | -27.6503
17.6291 | -34.9995 | -27.6577
17.6352 | -35 | -27.6653
17.6413 | -35.0005 | -27.6728
17.6474 | -35.001 | -27.6804
17.6535 | -35.0015 | -27.6878
17.6596 | -35.002 | -27.6951
17.6657 | -35.0025 | -27.7026
17.6718 | -35.003 | -27.71
17.6779 | -35.0035 | -27.7175
17.684 | -35.004 | -27.725
17.6901 | -35.0044 | -27.7323
17.6963 | -35.0049 | -27.7395
17.7024 | -35.0054 | -27.747
17.7085 | -35.0059 | -27.7543
17.7146 | -35.0064 | -27.7618
17.7207 | -35.0069 | -27.769
17.7268 | -35.0075 | -27.7763
17.7329 | -35.0079 | -27.7838
17.739 | -35.0084 | -27.7912
17.7451 | -35.0089 | -27.7986
17.7512 | -35.0093 | -27.8061
17.7573 | -35.0098 | -27.8136
17.7634 | -35.0103 | -27.8208
17.7695 | -35.0108 | -27.8282
17.7756 | -35.0113 | -27.8356
17.7817 | -35.0118 | -27.8428
17.7878 | -35.0123 | -27.8505
17.7939 | -35.0128 | -27.858
17.8 | -35.0133 | -27.8654
17.8061 | -35.0137 | -27.8727
17.8122 | -35.0143 | -27.8802
17.8183 | -35.0147 | -27.8876
17.8244 | -35.0152 | -27.8951
17.8305 | -35.0157 | -27.9025
17.8366 | -35.0162 | -27.9097
17.8428 | -35.0167 | -27.9171
17.8489 | -35.0172 | -27.9245
17.855 | -35.0177 | -27.9317
17.8611 | -35.0182 | -27.939
17.8672 | -35.0187 | -27.9463
17.8733 | -35.0192 | -27.9538
17.8794 | -35.0197 | -27.9611
17.8855 | -35.0202 | -27.9686
17.8916 | -35.0207 | -27.9763
17.8977 | -35.0211 | -27.9836
17.9038 | -35.0216 | -27.9906
17.9099 | -35.0221 | -27.9981
17.916 | -35.0226 | -28.0056
17.9221 | -35.0231 | -28.0129
17.9282 | -35.0236 | -28.0202
17.9343 | -35.0241 | -28.0278
17.9404 | -35.0246 | -28.035
17.9465 | -35.0252 | -28.0421
17.9526 | -35.0256 | -28.0497
17.9587 | -35.0261 | -28.0569
17.9648 | -35.0267 | -28.0641
17.9709 | -35.0272 | -28.0717
17.977 | -35.0276 | -28.0791
17.9832 | -35.0282 | -28.0865
17.9893 | -35.0287 | -28.0939
17.9954 | -35.0291 | -28.1012
18.0015 | -35.0296 | -28.1087
18.0076 | -35.0301 | -28.1158
18.0137 | -35.0306 | -28.1232
18.0198 | -35.0311 | -28.1303
18.0259 | -35.0316 | -28.1376
18.032 | -35.032 | -28.145
18.0381 | -35.0325 | -28.1524
18.0442 | -35.033 | -28.1597
18.0503 | -35.0335 | -28.1671
18.0564 | -35.034 | -28.1746
18.0625 | -35.0345 | -28.1821
18.0686 | -35.035 | -28.1895
18.0747 | -35.0355 | -28.1969
18.0808 | -35.0361 | -28.2041
18.0869 | -35.0365 | -28.2116
18.093 | -35.037 | -28.219
18.0991 | -35.0375 | -28.2262
18.1052 | -35.0379 | -28.2335
18.1113 | -35.0384 | -28.2409
18.1174 | -35.0389 | -28.2483
18.1236 | -35.0394 | -28.2557
18.1297 | -35.0399 | -28.2631
18.1358 | -35.0404 | -28.2705
18.1419 | -35.041 | -28.2777
18.148 | -35.0415 | -28.285
18.1541 | -35.042 | -28.2921
18.1602 | -35.0425 | -28.2996
18.1663 | -35.0429 | -28.3068
18.1724 | -35.0434 | -28.314
18.1785 | -35.0439 | -28.3214
18.1846 | -35.0443 | -28.3286
18.1907 | -35.0448 | -28.3359
18.1968 | -35.0453 | -28.3434
18.2029 | -35.0458 | -28.351
18.209 | -35.0463 | -28.3587
18.2151 | -35.0469 | -28.3663
18.2212 | -35.0474 | -28.3731
18.2273 | -35.0478 | -28.3804
18.2334 | -35.0483 | -28.3878
18.2395 | -35.0489 | -28.3952
18.2456 | -35.0494 | -28.4028
18.2517 | -35.0499 | -28.4101
18.2578 | -35.0504 | -28.4172
18.2639 | -35.0509 | -28.4245
18.2701 | -35.0514 | -28.4317
18.2762 | -35.0519 | -28.4391
18.2823 | -35.0524 | -28.4464
18.2884 | -35.0529 | -28.4536
18.2945 | -35.0535 | -28.4604
18.3006 | -35.054 | -28.4679
18.3067 | -35.0544 | -28.4751
18.3128 | -35.0549 | -28.4826
18.3189 | -35.0554 | -28.4898
18.325 | -35.0559 | -28.497
18.3311 | -35.0564 | -28.5043
18.3372 | -35.0569 | -28.5114
18.3433 | -35.0575 | -28.5189
18.3494 | -35.058 | -28.5262
18.3555 | -35.0585 | -28.5338
18.3616 | -35.059 | -28.541
18.3677 | -35.0595 | -28.5485
18.3738 | -35.06 | -28.5555
18.3799 | -35.0605 | -28.5628
18.386 | -35.061 | -28.5701
18.3921 | -35.0615 | -28.5775
18.3982 | -35.062 | -28.5847
18.4043 | -35.0625 | -28.5921
18.4105 | -35.063 | -28.5991
18.4166 | -35.0635 | -28.6064
18.4227 | -35.064 | -28.6135
18.4288 | -35.0645 | -28.6206
18.4349 | -35.065 | -28.6278
18.441 | -35.0655 | -28.6349
18.4471 | -35.066 | -28.642
18.4532 | -35.0665 | -28.6496
18.4593 | -35.0669 | -28.6567
18.4654 | -35.0674 | -28.6635
18.4715 | -35.0679 | -28.6711
18.4776 | -35.0685 | -28.6787
18.4837 | -35.0689 | -28.6858
18.4898 | -35.0695 | -28.6932
18.4959 | -35.07 | -28.7005
18.502 | -35.0704 | -28.7077
18.5081 | -35.071 | -28.7149
18.5142 | -35.0715 | -28.722
18.5203 | -35.072 | -28.7291
18.5264 | -35.0725 | -28.7364
18.5325 | -35.073 | -28.7438
18.5386 | -35.0735 | -28.751
18.5447 | -35.074 | -28.7583
18.5508 | -35.0745 | -28.7657
18.557 | -35.075 | -28.7732
18.5631 | -35.0755 | -28.7803
18.5692 | -35.076 | -28.7876
18.5753 | -35.0765 | -28.7946
18.5814 | -35.077 | -28.8017
18.5875 | -35.0775 | -28.8087
18.5936 | -35.078 | -28.8157
18.5997 | -35.0784 | -28.8231
18.6058 | -35.0789 | -28.83
18.6119 | -35.0794 | -28.8372
18.618 | -35.0799 | -28.8439
18.6241 | -35.0804 | -28.8509
18.6302 | -35.0809 | -28.8585
18.6363 | -35.0814 | -28.8658
18.6424 | -35.0819 | -28.8731
18.6485 | -35.0824 | -28.8802
18.6546 | -35.0829 | -28.8877
18.6607 | -35.0834 | -28.8949
18.6668 | -35.0839 | -28.9022
18.6729 | -35.0844 | -28.9096
18.679 | -35.0849 | -28.9169
18.6851 | -35.0854 | -28.9238
18.6912 | -35.086 | -28.9306
18.6974 | -35.0864 | -28.9379
18.7035 | -35.0869 | -28.9452
18.7096 | -35.0875 | -28.9529
18.7157 | -35.0879 | -28.9598
18.7218 | -35.0884 | -28.9669
18.7279 | -35.0888 | -28.974
18.734 | -35.0894 | -28.9814
18.7401 | -35.0899 | -28.9887
18.7462 | -35.0903 | -28.996
18.7523 | -35.0909 | -29.0037
18.7584 | -35.0914 | -29.0108
18.7645 | -35.0918 | -29.018
18.7706 | -35.0924 | -29.025
18.7767 | -35.0928 | -29.032
18.7828 | -35.0933 | -29.039
18.7889 | -35.0938 | -29.0464
18.795 | -35.0943 | -29.0538
18.8011 | -35.0948 | -29.0611
18.8072 | -35.0953 | -29.0683
18.8133 | -35.0958 | -29.0753
18.8194 | -35.0963 | -29.0824
18.8255 | -35.0968 | -29.0895
18.8316 | -35.0972 | -29.0968
18.8377 | -35.0978 | -29.104
18.8439 | -35.0983 | -29.1111
18.85 | -35.0988 | -29.1186
18.8561 | -35.0993 | -29.1259
18.8622 | -35.0998 | -29.133
18.8683 | -35.1002 | -29.1399
18.8744 | -35.1007 | -29.1469
18.8805 | -35.1012 | -29.1541
18.8866 | -35.1017 | -29.1614
18.8927 | -35.1022 | -29.1685
18.8988 | -35.1027 | -29.1758
18.9049 | -35.1032 | -29.1832
18.911 | -35.1037 | -29.1905
18.9171 | -35.1043 | -29.1979
18.9232 | -35.1047 | -29.2046
18.9293 | -35.1052 | -29.212
18.9354 | -35.1058 | -29.2193
18.9415 | -35.1063 | -29.2266
18.9476 | -35.1068 | -29.2334
18.9537 | -35.1073 | -29.2406
18.9598 | -35.1078 | -29.2474
18.9659 | -35.1083 | -29.2543
18.972 | -35.1088 | -29.2614
18.9781 | -35.1093 | -29.269
18.9843 | -35.1098 | -29.2761
18.9904 | -35.1103 | -29.2832
18.9965 | -35.1107 | -29.29
19.0026 | -35.1113 | -29.2973
19.0087 | -35.1118 | -29.3047
19.0148 | -35.1123 | -29.312
19.0209 | -35.1127 | -29.3193
19.027 | -35.1133 | -29.3267
19.0331 | -35.1137 | -29.3337
19.0392 | -35.1143 | -29.3408
19.0453 | -35.1148 | -29.3481
19.0514 | -35.1153 | -29.3552
19.0575 | -35.1158 | -29.3623
19.0636 | -35.1163 | -29.3692
19.0697 | -35.1168 | -29.3765
19.0758 | -35.1173 | -29.3835
19.0819 | -35.1178 | -29.3908
19.088 | -35.1184 | -29.3981
19.0941 | -35.1189 | -29.4051
19.1002 | -35.1194 | -29.4125
19.1063 | -35.1198 | -29.4196
19.1124 | -35.1203 | -29.4269
19.1185 | -35.1208 | -29.4343
19.1246 | -35.1213 | -29.4412
19.1308 | -35.1219 | -29.4484
19.1369 | -35.1223 | -29.4553
19.143 | -35.1228 | -29.4624
19.1491 | -35.1233 | -29.4694
19.1552 | -35.1238 | -29.4766
19.1613 | -35.1243 | -29.4838
19.1674 | -35.1248 | -29.4911
19.1735 | -35.1252 | -29.4985
19.1796 | -35.1256 | -29.5055
19.1857 | -35.126 | -29.5126
19.1918 | -35.1266 | -29.5196
19.1979 | -35.127 | -29.5271
19.204 | -35.1276 | -29.5341
19.2101 | -35.1282 | -29.5408
19.2162 | -35.1286 | -29.5481
19.2223 | -35.129 | -29.5551
19.2284 | -35.1295 | -29.5622
19.2345 | -35.13 | -29.57
19.2406 | -35.1305 | -29.5763
19.2467 | -35.131 | -29.5833
19.2528 | -35.1315 | -29.5905
19.2589 | -35.132 | -29.5972
19.265 | -35.1325 | -29.6041
19.2712 | -35.1329 | -29.611
19.2773 | -35.1334 | -29.6183
19.2834 | -35.1339 | -29.6253
19.2895 | -35.1344 | -29.6322
19.2956 | -35.1348 | -29.6395
19.3017 | -35.1353 | -29.6462
19.3078 | -35.1357 | -29.653
19.3139 | -35.1362 | -29.6601
19.32 | -35.1367 | -29.6674
19.3261 | -35.1372 | -29.6741
19.3322 | -35.1377 | -29.6809
19.3383 | -35.1382 | -29.6883
19.3444 | -35.1387 | -29.6956
19.3505 | -35.1391 | -29.7025
19.3566 | -35.1396 | -29.7098
19.3627 | -35.14 | -29.7172
19.3688 | -35.1405 | -29.7244
19.3749 | -35.141 | -29.7313
19.381 | -35.1414 | -29.7377
19.3871 | -35.1419 | -29.7449
19.3932 | -35.1424 | -29.7524
19.3993 | -35.1429 | -29.7594
19.4054 | -35.1433 | -29.7669
19.4115 | -35.1438 | -29.774
19.4177 | -35.1442 | -29.7807
19.4238 | -35.1447 | -29.7879
19.4299 | -35.1453 | -29.7952
19.436 | -35.1458 | -29.8024
19.4421 | -35.1462 | -29.8092
19.4482 | -35.1466 | -29.8159
19.4543 | -35.1471 | -29.8221
19.4604 | -35.1476 | -29.8298
19.4665 | -35.148 | -29.8368
19.4726 | -35.1485 | -29.8442
19.4787 | -35.1489 | -29.8516
19.4848 | -35.1494 | -29.859
19.4909 | -35.1498 | -29.8659
19.497 | -35.1504 | -29.8727
19.5031 | -35.1509 | -29.8797
19.5092 | -35.1513 | -29.8864
19.5153 | -35.1518 | -29.8934
19.5214 | -35.1523 | -29.8998
19.5275 | -35.1527 | -29.9068
19.5336 | -35.1532 | -29.9139
19.5397 | -35.1536 | -29.9208
19.5458 | -35.1541 | -29.9282
19.5519 | -35.1545 | -29.935
19.5581 | -35.1551 | -29.9422
19.5642 | -35.1555 | -29.9492
19.5703 | -35.156 | -29.956
19.5764 | -35.1564 | -29.9632
19.5825 | -35.1569 | -29.9703
19.5886 | -35.1574 | -29.9775
19.5947 | -35.1579 | -29.9841
19.6008 | -35.1584 | -29.9912
19.6069 | -35.1589 | -29.9981
19.613 | -35.1594 | -30.0054
19.6191 | -35.1598 | -30.0124
19.6252 | -35.1603 | -30.0194
19.6313 | -35.1607 | -30.0261
19.6374 | -35.1612 | -30.0332
19.6435 | -35.1617 | -30.04
19.6496 | -35.1622 | -30.0471
19.6557 | -35.1627 | -30.0542
19.6618 | -35.1631 | -30.0606
19.6679 | -35.1636 | -30.068
19.674 | -35.1641 | -30.0747
19.6801 | -35.1646 | -30.082
19.6862 | -35.1651 | -30.0888
19.6923 | -35.1656 | -30.0952
19.6984 | -35.1661 | -30.1019
19.7046 | -35.1665 | -30.1086
19.7107 | -35.167 | -30.1153
19.7168 | -35.1675 | -30.1221
19.7229 | -35.168 | -30.129
19.729 | -35.1685 | -30.1364
19.7351 | -35.169 | -30.1435
19.7412 | -35.1694 | -30.1508
19.7473 | -35.1699 | -30.158
19.7534 | -35.1704 | -30.1653
19.7595 | -35.1708 | -30.1729
19.7656 | -35.1713 | -30.1801
19.7717 | -35.1717 | -30.1871
19.7778 | -35.1721 | -30.194
19.7839 | -35.1726 | -30.2009
19.79 | -35.1731 | -30.2082
19.7961 | -35.1736 | -30.2154
19.8022 | -35.1741 | -30.2222
19.8083 | -35.1746 | -30.229
19.8144 | -35.175 | -30.2362
19.8205 | -35.1755 | -30.2429
19.8266 | -35.176 | -30.2501
19.8327 | -35.1765 | -30.2569
19.8388 | -35.1769 | -30.2638
19.845 | -35.1775 | -30.271
19.8511 | -35.1779 | -30.2779
19.8572 | -35.1783 | -30.2855
19.8633 | -35.1787 | -30.2922
19.8694 | -35.1792 | -30.299
19.8755 | -35.1797 | -30.3059
19.8816 | -35.1802 | -30.3129
19.8877 | -35.1806 | -30.3196
19.8938 | -35.1811 | -30.3262
19.8999 | -35.1815 | -30.3333
19.906 | -35.182 | -30.3401
19.9121 | -35.1826 | -30.3471
19.9182 | -35.1831 | -30.354
19.9243 | -35.1836 | -30.3606
19.9304 | -35.184 | -30.3674
19.9365 | -35.1845 | -30.3742
19.9426 | -35.1849 | -30.3811
19.9487 | -35.1854 | -30.388
19.9548 | -35.1858 | -30.3948
19.9609 | -35.1862 | -30.4018
19.967 | -35.1867 | -30.4087
19.9731 | -35.1872 | -30.4153
19.9792 | -35.1876 | -30.4219
19.9853 | -35.1881 | -30.4286
19.9915 | -35.1886 | -30.4357
19.9976 | -35.189 | -30.4422
20.0037 | -35.1894 | -30.4492
20.0098 | -35.1899 | -30.4565
20.0159 | -35.1904 | -30.4635
20.022 | -35.1908 | -30.4698
20.0281 | -35.1913 | -30.4762
20.0342 | -35.1917 | -30.4831
20.0403 | -35.1921 | -30.4902
20.0464 | -35.1926 | -30.497
20.0525 | -35.1931 | -30.5036
20.0586 | -35.1935 | -30.5104
20.0647 | -35.1939 | -30.5173
20.0708 | -35.1944 | -30.5243
20.0769 | -35.1948 | -30.5317
20.083 | -35.1953 | -30.5382
20.0891 | -35.1957 | -30.5454
20.0952 | -35.1961 | -30.552
20.1013 | -35.1965 | -30.5587
20.1074 | -35.1969 | -30.5654
20.1135 | -35.1974 | -30.5723
20.1196 | -35.1978 | -30.5792
20.1257 | -35.1983 | -30.5861
20.1319 | -35.1988 | -30.5928
20.138 | -35.1992 | -30.6
20.1441 | -35.1996 | -30.6068
20.1502 | -35.2001 | -30.6137
20.1563 | -35.2005 | -30.6201
20.1624 | -35.201 | -30.6266
20.1685 | -35.2015 | -30.6336
20.1746 | -35.202 | -30.6402
20.1807 | -35.2025 | -30.647
20.1868 | -35.203 | -30.6539
20.1929 | -35.2033 | -30.6609
20.199 | -35.2039 | -30.6678
20.2051 | -35.2044 | -30.6747
20.2112 | -35.2048 | -30.6815
20.2173 | -35.2053 | -30.6883
20.2234 | -35.2057 | -30.695
20.2295 | -35.2062 | -30.7013
20.2356 | -35.2066 | -30.7084
20.2417 | -35.207 | -30.7154
20.2478 | -35.2075 | -30.7222
20.2539 | -35.2079 | -30.7291
20.26 | -35.2083 | -30.7364
20.2661 | -35.2088 | -30.7434
20.2723 | -35.2092 | -30.7501
20.2784 | -35.2096 | -30.757
20.2845 | -35.2101 | -30.7637
20.2906 | -35.2105 | -30.7704
20.2967 | -35.2109 | -30.7771
20.3028 | -35.2114 | -30.7838
20.3089 | -35.2119 | -30.7907
20.315 | -35.2123 | -30.7978
20.3211 | -35.2128 | -30.8047
20.3272 | -35.2132 | -30.8122
20.3333 | -35.2137 | -30.8187
20.3394 | -35.2141 | -30.825
20.3455 | -35.2146 | -30.8319
20.3516 | -35.2151 | -30.8384
20.3577 | -35.2155 | -30.8453
20.3638 | -35.2159 | -30.8515
20.3699 | -35.2164 | -30.8578
20.376 | -35.2169 | -30.8649
20.3821 | -35.2174 | -30.8717
20.3882 | -35.2178 | -30.8787
20.3943 | -35.2182 | -30.8856
20.4004 | -35.2186 | -30.8919
20.4065 | -35.219 | -30.8984
20.4126 | -35.2195 | -30.9053
20.4188 | -35.22 | -30.9122
20.4249 | -35.2205 | -30.9187
20.431 | -35.2209 | -30.9255
20.4371 | -35.2213 | -30.9319
20.4432 | -35.2218 | -30.9392
20.4493 | -35.2222 | -30.9459
20.4554 | -35.2226 | -30.9532
20.4615 | -35.2231 | -30.96
20.4676 | -35.2235 | -30.9668
20.4737 | -35.2239 | -30.9733
20.4798 | -35.2243 | -30.9801
20.4859 | -35.2248 | -30.987
20.492 | -35.2252 | -30.9936
20.4981 | -35.2257 | -31.0003
20.5042 | -35.2261 | -31.007
20.5103 | -35.2265 | -31.0134
20.5164 | -35.227 | -31.02
20.5225 | -35.2274 | -31.0265
20.5286 | -35.2279 | -31.0331
20.5347 | -35.2283 | -31.0397
20.5408 | -35.2287 | -31.0463
20.5469 | -35.2292 | -31.0532
20.553 | -35.2296 | -31.0601
20.5592 | -35.23 | -31.067
20.5653 | -35.2304 | -31.0739
20.5714 | -35.2309 | -31.0806
20.5775 | -35.2314 | -31.0874
20.5836 | -35.2317 | -31.0944
20.5897 | -35.2322 | -31.1014
20.5958 | -35.2327 | -31.1082
20.6019 | -35.2331 | -31.1149
20.608 | -35.2335 | -31.1221
20.6141 | -35.234 | -31.1286
20.6202 | -35.2344 | -31.135
20.6263 | -35.2348 | -31.1415
20.6324 | -35.2353 | -31.1482
20.6385 | -35.2357 | -31.1551
20.6446 | -35.2362 | -31.1618
20.6507 | -35.2366 | -31.1681
20.6568 | -35.2371 | -31.1751
20.6629 | -35.2375 | -31.182
20.669 | -35.238 | -31.1888
20.6751 | -35.2384 | -31.1951
20.6812 | -35.2387 | -31.2017
20.6873 | -35.2392 | -31.2081
20.6934 | -35.2396 | -31.215
20.6995 | -35.24 | -31.2218
20.7057 | -35.2405 | -31.2288
20.7118 | -35.2409 | -31.2355
20.7179 | -35.2413 | -31.2421
20.724 | -35.2417 | -31.2489
20.7301 | -35.2421 | -31.2552
20.7362 | -35.2425 | -31.2622
20.7423 | -35.2429 | -31.2688
20.7484 | -35.2434 | -31.2759
20.7545 | -35.2438 | -31.2825
20.7606 | -35.2443 | -31.2893
20.7667 | -35.2447 | -31.2959
20.7728 | -35.2451 | -31.3027
20.7789 | -35.2455 | -31.309
20.785 | -35.246 | -31.3159
20.7911 | -35.2464 | -31.3226
20.7972 | -35.2468 | -31.3295
20.8033 | -35.2473 | -31.3369
20.8094 | -35.2477 | -31.3436
20.8155 | -35.2481 | -31.3507
20.8216 | -35.2486 | -31.3573
20.8277 | -35.2489 | -31.3636
20.8338 | -35.2493 | -31.3702
20.8399 | -35.2498 | -31.3769
20.8461 | -35.2502 | -31.3836
20.8522 | -35.2507 | -31.3905
20.8583 | -35.2511 | -31.3966
20.8644 | -35.2515 | -31.4033
20.8705 | -35.2519 | -31.4096
20.8766 | -35.2524 | -31.4164
20.8827 | -35.2528 | -31.4237
20.8888 | -35.2533 | -31.4303
20.8949 | -35.2537 | -31.4374
20.901 | -35.2541 | -31.4443
20.9071 | -35.2545 | -31.4508
20.9132 | -35.255 | -31.4573
20.9193 | -35.2554 | -31.4643
20.9254 | -35.2558 | -31.4706
20.9315 | -35.2563 | -31.4771
20.9376 | -35.2567 | -31.4842
20.9437 | -35.2572 | -31.4909
20.9498 | -35.2576 | -31.4976
20.9559 | -35.258 | -31.5045
20.962 | -35.2584 | -31.5109
20.9681 | -35.2588 | -31.5182
20.9742 | -35.2593 | -31.5248
20.9803 | -35.2598 | -31.5311
20.9864 | -35.2601 | -31.5377
20.9926 | -35.2607 | -31.5444
20.9987 | -35.2611 | -31.5514
21.0048 | -35.2615 | -31.5581
21.0109 | -35.2619 | -31.5647
21.017 | -35.2623 | -31.5714
21.0231 | -35.2628 | -31.5779
21.0292 | -35.2632 | -31.5846
21.0353 | -35.2636 | -31.5905
21.0414 | -35.264 | -31.5967
21.0475 | -35.2644 | -31.6034
21.0536 | -35.2648 | -31.61
21.0597 | -35.2653 | -31.6166
21.0658 | -35.2657 | -31.6234
21.0719 | -35.2662 | -31.6308
21.078 | -35.2666 | -31.6373
21.0841 | -35.267 | -31.6439
21.0902 | -35.2674 | -31.6504
21.0963 | -35.2677 | -31.6575
21.1024 | -35.268 | -31.6645
21.1085 | -35.2685 | -31.6706
21.1146 | -35.2689 | -31.6776
21.1207 | -35.2693 | -31.684
21.1268 | -35.2697 | -31.6906
21.133 | -35.2702 | -31.697
21.1391 | -35.2706 | -31.7036
21.1452 | -35.271 | -31.7105
21.1513 | -35.2715 | -31.717
21.1574 | -35.2719 | -31.7238
21.1635 | -35.2723 | -31.7302
21.1696 | -35.2727 | -31.7374
21.1757 | -35.2731 | -31.7436
21.1818 | -35.2735 | -31.7508
21.1879 | -35.274 | -31.7573
21.194 | -35.2744 | -31.7637
21.2001 | -35.2748 | -31.7704
21.2062 | -35.2752 | -31.7771
21.2123 | -35.2756 | -31.7835
21.2184 | -35.276 | -31.7895
21.2245 | -35.2764 | -31.7962
21.2306 | -35.2769 | -31.8026
21.2367 | -35.2773 | -31.8094
21.2428 | -35.2777 | -31.8161
21.2489 | -35.2781 | -31.8233
21.255 | -35.2785 | -31.8303
21.2611 | -35.279 | -31.8372
21.2672 | -35.2793 | -31.8436
21.2733 | -35.2796 | -31.8501
21.2795 | -35.2799 | -31.8565
21.2856 | -35.2803 | -31.8629
21.2917 | -35.2807 | -31.8698
21.2978 | -35.2811 | -31.8763
21.3039 | -35.2817 | -31.8828
21.31 | -35.2821 | -31.8898
21.3161 | -35.2826 | -31.8965
21.3222 | -35.283 | -31.9029
21.3283 | -35.2834 | -31.9099
21.3344 | -35.2837 | -31.9158
21.3405 | -35.2841 | -31.9225
21.3466 | -35.2845 | -31.9289
21.3527 | -35.2849 | -31.9358
21.3588 | -35.2852 | -31.9419
21.3649 | -35.2856 | -31.9492
21.371 | -35.2861 | -31.9557
21.3771 | -35.2865 | -31.9622
21.3832 | -35.2869 | -31.9688
21.3893 | -35.2871 | -31.976
21.3954 | -35.2876 | -31.9827
21.4015 | -35.288 | -31.9893
21.4076 | -35.2884 | -31.9962
21.4137 | -35.2888 | -32.0031
21.4199 | -35.2892 | -32.0102
21.426 | -35.2895 | -32.0162
21.4321 | -35.2899 | -32.023
21.4382 | -35.2903 | -32.0296
21.4443 | -35.2906 | -32.0365
21.4504 | -35.291 | -32.043
21.4565 | -35.2914 | -32.0498
21.4626 | -35.2918 | -32.0565
21.4687 | -35.2923 | -32.0634
21.4748 | -35.2927 | -32.0703
21.4809 | -35.2931 | -32.0768
21.487 | -35.2935 | -32.0832
21.4931 | -35.2939 | -32.0901
21.4992 | -35.2943 | -32.0961
21.5053 | -35.2947 | -32.1027
21.5114 | -35.2951 | -32.1098
21.5175 | -35.2955 | -32.1166
21.5236 | -35.2959 | -32.1231
21.5297 | -35.2962 | -32.1305
21.5358 | -35.2967 | -32.137
21.5419 | -35.297 | -32.1435
21.548 | -35.2974 | -32.15
21.5541 | -35.2978 | -32.1569
21.5602 | -35.2982 | -32.1637
21.5664 | -35.2985 | -32.1699
21.5725 | -35.2989 | -32.1761
21.5786 | -35.2994 | -32.1828
21.5847 | -35.2997 | -32.1897
21.5908 | -35.3001 | -32.1965
21.5969 | -35.3004 | -32.2027
21.603 | -35.3008 | -32.2089
21.6091 | -35.3012 | -32.2156
21.6152 | -35.3017 | -32.2221
21.6213 | -35.3021 | -32.2288
21.6274 | -35.3025 | -32.2359
21.6335 | -35.3028 | -32.2422
21.6396 | -35.3032 | -32.2483
21.6457 | -35.3037 | -32.2553
21.6518 | -35.3041 | -32.2613
21.6579 | -35.3045 | -32.2677
21.664 | -35.3049 | -32.2748
21.6701 | -35.3052 | -32.2825
21.6762 | -35.3056 | -32.2891
21.6823 | -35.306 | -32.2952
21.6884 | -35.3064 | -32.3025
21.6945 | -35.3069 | -32.3089
21.7006 | -35.3074 | -32.3155
21.7068 | -35.3079 | -32.3216
21.7129 | -35.3083 | -32.3279
21.719 | -35.3086 | -32.3345
21.7251 | -35.3089 | -32.3415
21.7312 | -35.3093 | -32.3476
21.7373 | -35.3097 | -32.3546
21.7434 | -35.31 | -32.3612
21.7495 | -35.3104 | -32.3682
21.7556 | -35.3108 | -32.3745
21.7617 | -35.3113 | -32.3813
21.7678 | -35.3117 | -32.3879
21.7739 | -35.312 | -32.3951
21.78 | -35.3124 | -32.4014
21.7861 | -35.3127 | -32.4082
21.7922 | -35.3131 | -32.4149
21.7983 | -35.3135 | -32.4217
21.8044 | -35.3139 | -32.4284
21.8105 | -35.3142 | -32.4354
21.8166 | -35.3147 | -32.4416
21.8227 | -35.315 | -32.4485
21.8288 | -35.3155 | -32.4555
21.8349 | -35.3158 | -32.4622
21.841 | -35.3162 | -32.4691
21.8471 | -35.3166 | -32.4758
21.8533 | -35.317 | -32.4822
21.8594 | -35.3175 | -32.4887
21.8655 | -35.3179 | -32.4948
21.8716 | -35.3183 | -32.501
21.8777 | -35.3187 | -32.5073
21.8838 | -35.3191 | -32.514
21.8899 | -35.3194 | -32.5201
21.896 | -35.3198 | -32.5275
21.9021 | -35.3202 | -32.5337
21.9082 | -35.3206 | -32.5409
21.9143 | -35.321 | -32.5469
21.9204 | -35.3214 | -32.5537
21.9265 | -35.3218 | -32.5606
21.9326 | -35.3221 | -32.5671
21.9387 | -35.3225 | -32.5739
21.9448 | -35.3229 | -32.5809
21.9509 | -35.3233 | -32.5876
21.957 | -35.3237 | -32.5942
21.9631 | -35.3241 | -32.6009
21.9692 | -35.3245 | -32.6077
21.9753 | -35.3248 | -32.6142
21.9814 | -35.3252 | -32.6214
21.9875 | -35.3256 | -32.6281
21.9937 | -35.3259 | -32.6351
21.9998 | -35.3263 | -32.6416
22.0059 | -35.3267 | -32.6482
22.012 | -35.327 | -32.6543
22.0181 | -35.3273 | -32.6601
22.0242 | -35.3277 | -32.6662
22.0303 | -35.3281 | -32.6731
22.0364 | -35.3285 | -32.6795
22.0425 | -35.3289 | -32.686
22.0486 | -35.3294 | -32.6928
22.0547 | -35.3297 | -32.6995
22.0608 | -35.3303 | -32.7058
22.0669 | -35.3306 | -32.7125
22.073 | -35.331 | -32.7189
22.0791 | -35.3314 | -32.7256
22.0852 | -35.3318 | -32.7323
22.0913 | -35.3323 | -32.7388
22.0974 | -35.3327 | -32.7454
22.1035 | -35.3331 | -32.7525
22.1096 | -35.3335 | -32.7589
22.1157 | -35.3339 | -32.7652
22.1218 | -35.3343 | -32.7718
22.1279 | -35.3346 | -32.7782
22.134 | -35.335 | -32.7846
22.1402 | -35.3354 | -32.7911
22.1463 | -35.3358 | -32.7974
22.1524 | -35.3362 | -32.8048
22.1585 | -35.3366 | -32.8115
22.1646 | -35.337 | -32.818
22.1707 | -35.3374 | -32.8251
22.1768 | -35.3378 | -32.8315
22.1829 | -35.3382 | -32.8387
22.189 | -35.3385 | -32.8453
22.1951 | -35.339 | -32.8519
22.2012 | -35.3394 | -32.8583
22.2073 | -35.3398 | -32.8649
22.2134 | -35.3402 | -32.8722
22.2195 | -35.3406 | -32.8792
22.2256 | -35.3409 | -32.8864
22.2317 | -35.3414 | -32.8926
22.2378 | -35.3418 | -32.8994
22.2439 | -35.3422 | -32.906
22.25 | -35.3425 | -32.9125
22.2561 | -35.3428 | -32.9192
22.2622 | -35.3433 | -32.9257
22.2683 | -35.3437 | -32.9326
22.2744 | -35.344 | -32.9394
22.2806 | -35.3444 | -32.9456
22.2867 | -35.3448 | -32.9527
22.2928 | -35.3452 | -32.9593
22.2989 | -35.3456 | -32.9661
22.305 | -35.346 | -32.9731
22.3111 | -35.3465 | -32.98
22.3172 | -35.3468 | -32.9865
22.3233 | -35.3472 | -32.9928
22.3294 | -35.3478 | -32.9997
22.3355 | -35.3483 | -33.0064
22.3416 | -35.3487 | -33.0129
22.3477 | -35.3491 | -33.0189
22.3538 | -35.3495 | -33.0251
22.3599 | -35.35 | -33.0315
22.366 | -35.3504 | -33.0385
22.3721 | -35.3509 | -33.045
22.3782 | -35.3514 | -33.0518
22.3843 | -35.3518 | -33.0585
22.3904 | -35.3521 | -33.065
22.3965 | -35.3524 | -33.0719
22.4026 | -35.3528 | -33.0781
22.4087 | -35.3532 | -33.0844
22.4148 | -35.3537 | -33.0907
22.4209 | -35.354 | -33.0967
22.4271 | -35.3544 | -33.1032
22.4332 | -35.3547 | -33.1097
22.4393 | -35.3551 | -33.1167
22.4454 | -35.3555 | -33.1232
22.4515 | -35.3559 | -33.1297
22.4576 | -35.3564 | -33.1367
22.4637 | -35.3567 | -33.1434
22.4698 | -35.3571 | -33.1495
22.4759 | -35.3576 | -33.1552
22.482 | -35.3579 | -33.162
22.4881 | -35.3583 | -33.1686
22.4942 | -35.3588 | -33.175
22.5003 | -35.359 | -33.182
22.5064 | -35.3594 | -33.1889
22.5125 | -35.3598 | -33.1957
22.5186 | -35.3602 | -33.2021
22.5247 | -35.3606 | -33.2084
22.5308 | -35.361 | -33.2143
22.5369 | -35.3613 | -33.2214
22.543 | -35.3617 | -33.2278
22.5491 | -35.3622 | -33.2339
22.5552 | -35.3626 | -33.2402
22.5613 | -35.363 | -33.2473
22.5675 | -35.3634 | -33.2533
22.5736 | -35.3638 | -33.2606
22.5797 | -35.3642 | -33.2668
22.5858 | -35.3646 | -33.2731
22.5919 | -35.3649 | -33.2797
22.598 | -35.3654 | -33.2869
22.6041 | -35.3658 | -33.2932
22.6102 | -35.3661 | -33.2997
22.6163 | -35.3666 | -33.3059
22.6224 | -35.367 | -33.3122
22.6285 | -35.3674 | -33.3187
22.6346 | -35.3678 | -33.3255
22.6407 | -35.3683 | -33.3321
22.6468 | -35.3687 | -33.3386
22.6529 | -35.3692 | -33.3459
22.659 | -35.3695 | -33.3524
22.6651 | -35.37 | -33.3594
22.6712 | -35.3704 | -33.3663
22.6773 | -35.3708 | -33.3734
22.6834 | -35.3712 | -33.3801
22.6895 | -35.3716 | -33.386
22.6956 | -35.372 | -33.3929
22.7017 | -35.3724 | -33.3989
22.7079 | -35.3729 | -33.406
22.714 | -35.3733 | -33.413
22.7201 | -35.3737 | -33.4199
22.7262 | -35.374 | -33.4271
22.7323 | -35.3744 | -33.4336
22.7384 | -35.3748 | -33.4402
22.7445 | -35.3751 | -33.4465
22.7506 | -35.3755 | -33.4532
22.7567 | -35.3758 | -33.4603
22.7628 | -35.3762 | -33.467
22.7689 | -35.3766 | -33.4734
22.775 | -35.377 | -33.4805
22.7811 | -35.3775 | -33.4872
22.7872 | -35.3779 | -33.4935
22.7933 | -35.3783 | -33.5003
22.7994 | -35.3787 | -33.5063
22.8055 | -35.379 | -33.5129
22.8116 | -35.3795 | -33.5195
22.8177 | -35.3798 | -33.526
22.8238 | -35.3801 | -33.5326
22.8299 | -35.3806 | -33.5387
22.836 | -35.381 | -33.5453
22.8421 | -35.3815 | -33.552
22.8482 | -35.3819 | -33.5587
22.8544 | -35.3823 | -33.565
22.8605 | -35.3828 | -33.5718
22.8666 | -35.3832 | -33.5787
22.8727 | -35.3836 | -33.5855
22.8788 | -35.384 | -33.5916
22.8849 | -35.3845 | -33.5981
22.891 | -35.3849 | -33.6046
22.8971 | -35.3854 | -33.6116
22.9032 | -35.3857 | -33.6185
22.9093 | -35.3862 | -33.625
22.9154 | -35.3866 | -33.6309
22.9215 | -35.3871 | -33.6372
22.9276 | -35.3875 | -33.644
22.9337 | -35.3879 | -33.6508
22.9398 | -35.3883 | -33.6577
22.9459 | -35.3887 | -33.664
22.952 | -35.3891 | -33.6707
22.9581 | -35.3895 | -33.6778
22.9642 | -35.3898 | -33.6846
22.9703 | -35.3902 | -33.6911
22.9764 | -35.3906 | -33.6974
22.9825 | -35.3911 | -33.7038
22.9886 | -35.3915 | -33.7108
22.9948 | -35.3919 | -33.7174
23.0009 | -35.3922 | -33.7238
23.007 | -35.3926 | -33.7304
23.0131 | -35.393 | -33.7373
23.0192 | -35.3933 | -33.7433
23.0253 | -35.3937 | -33.7497
23.0314 | -35.3941 | -33.7568
23.0375 | -35.3945 | -33.7635
23.0436 | -35.3948 | -33.7699
23.0497 | -35.3952 | -33.7765
23.0558 | -35.3955 | -33.7831
23.0619 | -35.396 | -33.7901
23.068 | -35.3964 | -33.7969
23.0741 | -35.3968 | -33.8038
23.0802 | -35.3971 | -33.8103
23.0863 | -35.3974 | -33.8166
23.0924 | -35.3978 | -33.8232
23.0985 | -35.3982 | -33.8296
23.1046 | -35.3986 | -33.8362
23.1107 | -35.399 | -33.8433
23.1168 | -35.3995 | -33.8499
23.1229 | -35.3999 | -33.8562
23.129 | -35.4004 | -33.863
23.1351 | -35.4007 | -33.8698
23.1413 | -35.4011 | -33.8761
23.1474 | -35.4016 | -33.8824
23.1535 | -35.4019 | -33.8892
23.1596 | -35.4024 | -33.8953
23.1657 | -35.4028 | -33.9021
23.1718 | -35.4032 | -33.9081
23.1779 | -35.4035 | -33.915
23.184 | -35.404 | -33.9214
23.1901 | -35.4043 | -33.9279
23.1962 | -35.4047 | -33.9347
23.2023 | -35.4052 | -33.9411
23.2084 | -35.4056 | -33.9473
23.2145 | -35.4061 | -33.9533
23.2206 | -35.4065 | -33.9601
23.2267 | -35.4069 | -33.9668
23.2328 | -35.4073 | -33.9732
23.2389 | -35.4077 | -33.9802
23.245 | -35.4082 | -33.9867
23.2511 | -35.4087 | -33.9932
23.2572 | -35.4091 | -34.0004
23.2633 | -35.4095 | -34.0066
23.2694 | -35.4098 | -34.0135
23.2755 | -35.4102 | -34.0202
23.2817 | -35.4107 | -34.0264
23.2878 | -35.411 | -34.0328
23.2939 | -35.4115 | -34.0395
23.3 | -35.4119 | -34.0459
23.3061 | -35.4124 | -34.0525
23.3122 | -35.4128 | -34.0588
23.3183 | -35.4133 | -34.065
23.3244 | -35.4137 | -34.0717
23.3305 | -35.4141 | -34.0786
23.3366 | -35.4146 | -34.0849
23.3427 | -35.415 | -34.0912
23.3488 | -35.4154 | -34.0981
23.3549 | -35.4158 | -34.1046
23.361 | -35.4163 | -34.1112
23.3671 | -35.4167 | -34.1175
23.3732 | -35.4171 | -34.1242
23.3793 | -35.4175 | -34.1309
23.3854 | -35.4179 | -34.1376
23.3915 | -35.4183 | -34.1437
23.3976 | -35.4188 | -34.1507
23.4037 | -35.4192 | -34.1573
23.4098 | -35.4197 | -34.1639
23.4159 | -35.42 | -34.1704
23.422 | -35.4205 | -34.1767
23.4282 | -35.4209 | -34.1833
23.4343 | -35.4213 | -34.1899
23.4404 | -35.4217 | -34.1964
23.4465 | -35.4221 | -34.2031
23.4526 | -35.4225 | -34.2098
23.4587 | -35.4229 | -34.2163
23.4648 | -35.4233 | -34.2225
23.4709 | -35.4237 | -34.2292
23.477 | -35.4241 | -34.2356
23.4831 | -35.4245 | -34.2419
23.4892 | -35.4249 | -34.2484
23.4953 | -35.4253 | -34.2554
23.5014 | -35.4257 | -34.2619
23.5075 | -35.4261 | -34.2688
23.5136 | -35.4266 | -34.2755
23.5197 | -35.427 | -34.2818
23.5258 | -35.4275 | -34.2882
23.5319 | -35.4279 | -34.2949
23.538 | -35.4283 | -34.3017
23.5441 | -35.4286 | -34.3081
23.5502 | -35.429 | -34.3145
23.5563 | -35.4294 | -34.3212
23.5624 | -35.4298 | -34.3275
23.5686 | -35.4302 | -34.3338
23.5747 | -35.4306 | -34.3402
23.5808 | -35.4311 | -34.3467
23.5869 | -35.4315 | -34.353
23.593 | -35.4319 | -34.3596
23.5991 | -35.4324 | -34.3661
23.6052 | -35.4328 | -34.373
23.6113 | -35.4333 | -34.3796
23.6174 | -35.4337 | -34.3862
23.6235 | -35.4341 | -34.3925
23.6296 | -35.4345 | -34.3988
23.6357 | -35.435 | -34.4053
23.6418 | -35.4354 | -34.412
23.6479 | -35.4358 | -34.4181
23.654 | -35.4362 | -34.4245
23.6601 | -35.4366 | -34.4311
23.6662 | -35.437 | -34.4373
23.6723 | -35.4374 | -34.4436
23.6784 | -35.4378 | -34.4502
23.6845 | -35.4382 | -34.4564
23.6906 | -35.4387 | -34.4629
23.6967 | -35.439 | -34.4695
23.7028 | -35.4395 | -34.4762
23.7089 | -35.4399 | -34.4822
23.7151 | -35.4403 | -34.4886
23.7212 | -35.4408 | -34.4953
23.7273 | -35.4411 | -34.5018
23.7334 | -35.4416 | -34.508
23.7395 | -35.442 | -34.5142
23.7456 | -35.4425 | -34.5208
23.7517 | -35.4429 | -34.5271
23.7578 | -35.4434 | -34.5339
23.7639 | -35.4438 | -34.5405
23.77 | -35.4442 | -34.5469
23.7761 | -35.4446 | -34.5536
23.7822 | -35.445 | -34.5603
23.7883 | -35.4454 | -34.5665
23.7944 | -35.4459 | -34.5731
23.8005 | -35.4463 | -34.5794
23.8066 | -35.4467 | -34.5858
23.8127 | -35.4472 | -34.5921
23.8188 | -35.4476 | -34.5985
23.8249 | -35.448 | -34.6048
23.831 | -35.4485 | -34.6114
23.8371 | -35.4489 | -34.6179
23.8432 | -35.4493 | -34.6239
23.8493 | -35.4498 | -34.6303
23.8555 | -35.4502 | -34.6368
23.8616 | -35.4506 | -34.6434
23.8677 | -35.451 | -34.6497
23.8738 | -35.4515 | -34.6559
23.8799 | -35.4519 | -34.6625
23.886 | -35.4523 | -34.6685
23.8921 | -35.4527 | -34.675
23.8982 | -35.4531 | -34.6814
23.9043 | -35.4535 | -34.6875
23.9104 | -35.4539 | -34.6937
23.9165 | -35.4543 | -34.7004
23.9226 | -35.4547 | -34.7068
23.9287 | -35.4551 | -34.7134
23.9348 | -35.4555 | -34.7201
23.9409 | -35.4559 | -34.7259
23.947 | -35.4564 | -34.7321
23.9531 | -35.4568 | -34.7383
23.9592 | -35.4572 | -34.745
23.9653 | -35.4575 | -34.7511
23.9714 | -35.4579 | -34.7572
23.9775 | -35.4584 | -34.763
23.9836 | -35.4589 | -34.7696
23.9897 | -35.4593 | -34.7758
23.9958 | -35.4597 | -34.7825
24.002 | -35.4601 | -34.7889
24.0081 | -35.4605 | -34.7951
24.0142 | -35.4609 | -34.8021
24.0203 | -35.4614 | -34.8087
24.0264 | -35.4618 | -34.8149
24.0325 | -35.4622 | -34.8207
24.0386 | -35.4626 | -34.827
24.0447 | -35.463 | -34.8338
24.0508 | -35.4635 | -34.8395
24.0569 | -35.4638 | -34.8458
24.063 | -35.4644 | -34.8528
24.0691 | -35.4648 | -34.8594
24.0752 | -35.4652 | -34.8659
24.0813 | -35.4657 | -34.8725
24.0874 | -35.4662 | -34.8788
24.0935 | -35.4667 | -34.8853
24.0996 | -35.467 | -34.8914
24.1057 | -35.4675 | -34.8983
24.1118 | -35.4679 | -34.9048
24.1179 | -35.4683 | -34.9112
24.124 | -35.4687 | -34.9177
24.1301 | -35.4692 | -34.924
24.1362 | -35.4696 | -34.9305
24.1424 | -35.47 | -34.9369
24.1485 | -35.4704 | -34.9437
24.1546 | -35.4709 | -34.9499
24.1607 | -35.4713 | -34.956
24.1668 | -35.4717 | -34.9625
24.1729 | -35.4721 | -34.9687
24.179 | -35.4725 | -34.9751
24.1851 | -35.473 | -34.9813
24.1912 | -35.4734 | -34.9884
24.1973 | -35.4738 | -34.9948
24.2034 | -35.4742 | -35.0003
24.2095 | -35.4746 | -35.0065
24.2156 | -35.4749 | -35.0128
24.2217 | -35.4754 | -35.0193
24.2278 | -35.4759 | -35.0256
24.2339 | -35.4764 | -35.032
24.24 | -35.4768 | -35.0381
24.2461 | -35.4773 | -35.0448
24.2522 | -35.4777 | -35.0518
24.2583 | -35.4782 | -35.058
24.2644 | -35.4786 | -35.0644
24.2705 | -35.4791 | -35.0712
24.2766 | -35.4795 | -35.0775
24.2827 | -35.48 | -35.0838
24.2889 | -35.4804 | -35.0906
24.295 | -35.4808 | -35.0967
24.3011 | -35.4812 | -35.1024
24.3072 | -35.4818 | -35.1082
24.3133 | -35.4822 | -35.1146
24.3194 | -35.4826 | -35.1204
24.3255 | -35.483 | -35.1267
24.3316 | -35.4834 | -35.1332
24.3377 | -35.4839 | -35.1392
24.3438 | -35.4844 | -35.1449
24.3499 | -35.4847 | -35.1515
24.356 | -35.4852 | -35.1581
24.3621 | -35.4857 | -35.1643
24.3682 | -35.4861 | -35.1705
24.3743 | -35.4866 | -35.1771
24.3804 | -35.4871 | -35.1835
24.3865 | -35.4875 | -35.1894
24.3926 | -35.4879 | -35.196
24.3987 | -35.4883 | -35.2025
24.4048 | -35.4888 | -35.2093
24.4109 | -35.4892 | -35.2153
24.417 | -35.4897 | -35.2214
24.4231 | -35.4902 | -35.2275
24.4293 | -35.4906 | -35.2338
24.4354 | -35.4911 | -35.2399
24.4415 | -35.4915 | -35.2462
24.4476 | -35.492 | -35.2526
24.4537 | -35.4925 | -35.2588
24.4598 | -35.4929 | -35.2651
24.4659 | -35.4934 | -35.2717
24.472 | -35.4939 | -35.2775
24.4781 | -35.4944 | -35.2842
24.4842 | -35.4948 | -35.2898
24.4903 | -35.4953 | -35.2966
24.4964 | -35.4957 | -35.3029
24.5025 | -35.4962 | -35.3093
24.5086 | -35.4966 | -35.3155
24.5147 | -35.4971 | -35.3211
24.5208 | -35.4976 | -35.3273
24.5269 | -35.4981 | -35.3337
24.533 | -35.4986 | -35.3398
24.5391 | -35.499 | -35.3459
24.5452 | -35.4995 | -35.3522
24.5513 | -35.5 | -35.3571
24.5574 | -35.5005 | -35.3632
24.5635 | -35.5009 | -35.3695
24.5696 | -35.5013 | -35.376
24.5758 | -35.5017 | -35.3822
24.5819 | -35.5022 | -35.3884
24.588 | -35.5026 | -35.3941
24.5941 | -35.503 | -35.4001
24.6002 | -35.5034 | -35.4062
24.6063 | -35.5039 | -35.4122
24.6124 | -35.5043 | -35.418
24.6185 | -35.5048 | -35.4243
24.6246 | -35.5053 | -35.4305
24.6307 | -35.5057 | -35.4368
24.6368 | -35.5062 | -35.4431
24.6429 | -35.5067 | -35.4496
24.649 | -35.5072 | -35.4552
24.6551 | -35.5077 | -35.4614
24.6612 | -35.5081 | -35.4668
24.6673 | -35.5085 | -35.4736
24.6734 | -35.509 | -35.48
24.6795 | -35.5094 | -35.486
24.6856 | -35.5099 | -35.4921
24.6917 | -35.5104 | -35.4983
24.6978 | -35.5108 | -35.5043
24.7039 | -35.5113 | -35.5114
24.71 | -35.5118 | -35.5176
24.7162 | -35.5122 | -35.524
24.7223 | -35.5127 | -35.5302
24.7284 | -35.5132 | -35.5354
24.7345 | -35.5137 | -35.5415
24.7406 | -35.5141 | -35.5475
24.7467 | -35.5146 | -35.5535
24.7528 | -35.515 | -35.5594
24.7589 | -35.5155 | -35.5652
24.765 | -35.5159 | -35.5707
24.7711 | -35.5163 | -35.577
24.7772 | -35.5168 | -35.5826
24.7833 | -35.5173 | -35.5886
24.7894 | -35.5178 | -35.5952
24.7955 | -35.5183 | -35.6017
24.8016 | -35.5188 | -35.6072
24.8077 | -35.5192 | -35.6133
24.8138 | -35.5196 | -35.6203
24.8199 | -35.52 | -35.6262
24.826 | -35.5203 | -35.632
24.8321 | -35.5208 | -35.6383
24.8382 | -35.5211 | -35.6444
24.8443 | -35.5215 | -35.6508
24.8504 | -35.5219 | -35.6564
24.8565 | -35.5224 | -35.6625
24.8627 | -35.5229 | -35.6684
24.8688 | -35.5234 | -35.6737
24.8749 | -35.5238 | -35.6793
24.881 | -35.5242 | -35.6859
24.8871 | -35.5247 | -35.6921
24.8932 | -35.5251 | -35.6979
24.8993 | -35.5255 | -35.7042
24.9054 | -35.5259 | -35.7096
24.9115 | -35.5264 | -35.7162
24.9176 | -35.5268 | -35.7217
24.9237 | -35.5273 | -35.7281
24.9298 | -35.5277 | -35.7339
24.9359 | -35.5282 | -35.7402
24.942 | -35.5287 | -35.7455
24.9481 | -35.5291 | -35.7517
24.9542 | -35.5296 | -35.758
24.9603 | -35.53 | -35.7642
24.9664 | -35.5305 | -35.7698
24.9725 | -35.5309 | -35.7757
24.9786 | -35.5314 | -35.7812
24.9847 | -35.5319 | -35.7871
24.9908 | -35.5323 | -35.7932
24.9969 | -35.5328 | -35.7997
25.0031 | -35.5332 | -35.8054
25.0092 | -35.5336 | -35.8106
25.0153 | -35.534 | -35.8167
25.0214 | -35.5345 | -35.8227
25.0275 | -35.5349 | -35.8286
25.0336 | -35.5353 | -35.8345
25.0397 | -35.5357 | -35.8403
25.0458 | -35.5362 | -35.846
25.0519 | -35.5366 | -35.852
25.058 | -35.5371 | -35.8576
25.0641 | -35.5374 | -35.8634
25.0702 | -35.5378 | -35.8693
25.0763 | -35.5382 | -35.8752
25.0824 | -35.5387 | -35.8812
25.0885 | -35.5391 | -35.8875
25.0946 | -35.5396 | -35.8932
25.1007 | -35.54 | -35.8991
25.1068 | -35.5405 | -35.905
25.1129 | -35.5409 | -35.911
25.119 | -35.5414 | -35.9171
25.1251 | -35.5418 | -35.9232
25.1312 | -35.5422 | -35.9293
25.1373 | -35.5427 | -35.9351
25.1435 | -35.5432 | -35.941
25.1496 | -35.5435 | -35.9468
25.1557 | -35.5438 | -35.9525
25.1618 | -35.5444 | -35.9588
25.1679 | -35.5448 | -35.9649
25.174 | -35.5452 | -35.9708
25.1801 | -35.5456 | -35.9771
25.1862 | -35.5461 | -35.9826
25.1923 | -35.5465 | -35.9888
25.1984 | -35.547 | -35.9953
25.2045 | -35.5475 | -36.0016
25.2106 | -35.5479 | -36.007
25.2167 | -35.5484 | -36.0129
25.2228 | -35.5488 | -36.018
25.2289 | -35.5493 | -36.0236
25.235 | -35.5497 | -36.0295
25.2411 | -35.5503 | -36.0353
25.2472 | -35.5507 | -36.0411
25.2533 | -35.5512 | -36.047
25.2594 | -35.5516 | -36.0526
25.2655 | -35.552 | -36.0588
25.2716 | -35.5525 | -36.0644
25.2777 | -35.5529 | -36.0701
25.2838 | -35.5533 | -36.0757
25.29 | -35.5538 | -36.0819
25.2961 | -35.5543 | -36.0876
25.3022 | -35.5547 | -36.0934
25.3083 | -35.5551 | -36.0992
25.3144 | -35.5556 | -36.1055
25.3205 | -35.556 | -36.111
25.3266 | -35.5565 | -36.1167
25.3327 | -35.5569 | -36.1221
25.3388 | -35.5574 | -36.1278
25.3449 | -35.5579 | -36.1337
25.351 | -35.5583 | -36.139
25.3571 | -35.5587 | -36.1445
25.3632 | -35.5592 | -36.15
25.3693 | -35.5597 | -36.1561
25.3754 | -35.5601 | -36.1622
25.3815 | -35.5605 | -36.1678
25.3876 | -35.561 | -36.1741
25.3937 | -35.5615 | -36.1802
25.3998 | -35.562 | -36.1852
25.4059 | -35.5624 | -36.1911
25.412 | -35.5628 | -36.1971
25.4181 | -35.5632 | -36.2026
25.4242 | -35.5636 | -36.2081
25.4304 | -35.5641 | -36.2141
25.4365 | -35.5645 | -36.2198
25.4426 | -35.565 | -36.2255
25.4487 | -35.5654 | -36.2311
25.4548 | -35.566 | -36.2371
25.4609 | -35.5665 | -36.2429
25.467 | -35.5669 | -36.2481
25.4731 | -35.5674 | -36.2545
25.4792 | -35.5679 | -36.2596
25.4853 | -35.5683 | -36.2654
25.4914 | -35.5688 | -36.2716
25.4975 | -35.5692 | -36.2773
25.5036 | -35.5696 | -36.283
25.5097 | -35.5702 | -36.2884
25.5158 | -35.5706 | -36.2947
25.5219 | -35.571 | -36.3006
25.528 | -35.5714 | -36.3064
25.5341 | -35.5719 | -36.3118
25.5402 | -35.5723 | -36.3166
25.5463 | -35.5728 | -36.3225
25.5524 | -35.5732 | -36.3281
25.5585 | -35.5736 | -36.3344
25.5646 | -35.5739 | -36.3399
25.5707 | -35.5744 | -36.3454
25.5769 | -35.5748 | -36.3505
25.583 | -35.5752 | -36.3562
25.5891 | -35.5757 | -36.3618
25.5952 | -35.5761 | -36.3673
25.6013 | -35.5765 | -36.3728
25.6074 | -35.5769 | -36.3787
25.6135 | -35.5774 | -36.3842
25.6196 | -35.5779 | -36.3898
25.6257 | -35.5783 | -36.3956
25.6318 | -35.5787 | -36.401
25.6379 | -35.5793 | -36.4069
25.644 | -35.5797 | -36.4118
25.6501 | -35.58 | -36.4171
25.6562 | -35.5805 | -36.4225
25.6623 | -35.581 | -36.4282
25.6684 | -35.5814 | -36.434
25.6745 | -35.5819 | -36.4402
25.6806 | -35.5824 | -36.4453
25.6867 | -35.5829 | -36.4512
25.6928 | -35.5833 | -36.4574
25.6989 | -35.5838 | -36.4628
25.705 | -35.5843 | -36.4682
25.7111 | -35.5847 | -36.4741
25.7173 | -35.5853 | -36.4795
25.7234 | -35.5856 | -36.4853
25.7295 | -35.5861 | -36.491
25.7356 | -35.5865 | -36.4967
25.7417 | -35.587 | -36.5026
25.7478 | -35.5874 | -36.5077
25.7539 | -35.5878 | -36.5128
25.76 | -35.5881 | -36.5181
25.7661 | -35.5885 | -36.5238
25.7722 | -35.589 | -36.5296
25.7783 | -35.5895 | -36.5347
25.7844 | -35.5899 | -36.5393
25.7905 | -35.5903 | -36.5453
25.7966 | -35.5908 | -36.5516
25.8027 | -35.5911 | -36.5571
25.8088 | -35.5916 | -36.5635
25.8149 | -35.5921 | -36.5696
25.821 | -35.5923 | -36.5744
25.8271 | -35.5928 | -36.58
25.8332 | -35.5932 | -36.5855
25.8393 | -35.5937 | -36.5915
25.8454 | -35.5942 | -36.5971
25.8515 | -35.5946 | -36.6029
25.8576 | -35.595 | -36.6084
25.8638 | -35.5955 | -36.614
25.8699 | -35.5959 | -36.6192
25.876 | -35.5963 | -36.6244
25.8821 | -35.5968 | -36.63
25.8882 | -35.5972 | -36.6348
25.8943 | -35.5977 | -36.6401
25.9004 | -35.5981 | -36.6457
25.9065 | -35.5985 | -36.6506
25.9126 | -35.5989 | -36.6559
25.9187 | -35.5993 | -36.6615
25.9248 | -35.5996 | -36.6669
25.9309 | -35.6 | -36.673
25.937 | -35.6004 | -36.6789
25.9431 | -35.6007 | -36.6847
25.9492 | -35.6012 | -36.6902
25.9553 | -35.6016 | -36.6957
25.9614 | -35.6021 | -36.7005
25.9675 | -35.6025 | -36.7061
25.9736 | -35.6029 | -36.7115
25.9797 | -35.6033 | -36.717
25.9858 | -35.6037 | -36.7229
25.9919 | -35.6041 | -36.7286
25.998 | -35.6044 | -36.7346
26.0042 | -35.6048 | -36.7399
26.0103 | -35.6053 | -36.7455
26.0164 | -35.6057 | -36.7505
26.0225 | -35.6062 | -36.756
26.0286 | -35.6066 | -36.7612
26.0347 | -35.607 | -36.766
26.0408 | -35.6075 | -36.7719
26.0469 | -35.6078 | -36.7774
26.053 | -35.6083 | -36.7836
26.0591 | -35.6088 | -36.7888
26.0652 | -35.6091 | -36.7939
26.0713 | -35.6096 | -36.7994
26.0774 | -35.61 | -36.8048
26.0835 | -35.6105 | -36.8103
26.0896 | -35.6109 | -36.8162
26.0957 | -35.6114 | -36.8213
26.1018 | -35.6119 | -36.8274
26.1079 | -35.6123 | -36.8328
26.114 | -35.6128 | -36.8388
26.1201 | -35.6133 | -36.8441
26.1262 | -35.6136 | -36.8496
26.1323 | -35.614 | -36.8552
26.1384 | -35.6143 | -36.8611
26.1445 | -35.6146 | -36.8668
26.1507 | -35.615 | -36.8725
26.1568 | -35.6154 | -36.8785
26.1629 | -35.6157 | -36.8839
26.169 | -35.6162 | -36.8891
26.1751 | -35.6167 | -36.894
26.1812 | -35.617 | -36.8998
26.1873 | -35.6176 | -36.9055
26.1934 | -35.618 | -36.9109
26.1995 | -35.6185 | -36.9164
26.2056 | -35.6187 | -36.921
26.2117 | -35.6192 | -36.9264
26.2178 | -35.6197 | -36.9313
26.2239 | -35.6201 | -36.9368
26.23 | -35.6205 | -36.9429
26.2361 | -35.6209 | -36.948
26.2422 | -35.6214 | -36.9532
26.2483 | -35.6217 | -36.9586
26.2544 | -35.622 | -36.9642
26.2605 | -35.6224 | -36.9697
26.2666 | -35.623 | -36.9759
26.2727 | -35.6235 | -36.9817
26.2788 | -35.6241 | -36.9867
26.2849 | -35.6245 | -36.9922
26.2911 | -35.625 | -36.9972
26.2972 | -35.6255 | -37.0027
26.3033 | -35.6259 | -37.0077
26.3094 | -35.6263 | -37.0136
26.3155 | -35.6266 | -37.0182
26.3216 | -35.6269 | -37.0243
26.3277 | -35.6274 | -37.0302
26.3338 | -35.6278 | -37.0352
26.3399 | -35.6282 | -37.0406
26.346 | -35.6287 | -37.0441
26.3521 | -35.6292 | -37.0492
26.3582 | -35.6296 | -37.0546
26.3643 | -35.63 | -37.0602
26.3704 | -35.6304 | -37.0654
26.3765 | -35.6309 | -37.0712
26.3826 | -35.6312 | -37.076
26.3887 | -35.6316 | -37.0811
26.3948 | -35.6321 | -37.0867
26.4009 | -35.6325 | -37.0925
26.407 | -35.6329 | -37.0986
26.4131 | -35.6333 | -37.1044
26.4192 | -35.6336 | -37.1095
26.4253 | -35.634 | -37.1147
26.4314 | -35.6344 | -37.1203
26.4376 | -35.6349 | -37.1252
26.4437 | -35.6354 | -37.1303
26.4498 | -35.6358 | -37.1362
26.4559 | -35.6363 | -37.141
26.462 | -35.6367 | -37.1465
26.4681 | -35.637 | -37.152
26.4742 | -35.6375 | -37.1573
26.4803 | -35.6379 | -37.1625
26.4864 | -35.6384 | -37.1671
26.4925 | -35.6387 | -37.1723
26.4986 | -35.6391 | -37.1775
26.5047 | -35.6396 | -37.1823
26.5108 | -35.64 | -37.1883
26.5169 | -35.6404 | -37.1936
26.523 | -35.6409 | -37.1986
26.5291 | -35.6412 | -37.204
26.5352 | -35.6417 | -37.2096
26.5413 | -35.6421 | -37.2147
26.5474 | -35.6424 | -37.2203
26.5535 | -35.6428 | -37.2268
26.5596 | -35.6432 | -37.2314
26.5657 | -35.6437 | -37.2364
26.5718 | -35.6442 | -37.2415
26.578 | -35.6447 | -37.2469
26.5841 | -35.6452 | -37.2523
26.5902 | -35.6456 | -37.2585
26.5963 | -35.646 | -37.2646
26.6024 | -35.6465 | -37.2706
26.6085 | -35.6469 | -37.276
26.6146 | -35.6474 | -37.2811
26.6207 | -35.6479 | -37.2867
26.6268 | -35.6483 | -37.2915
26.6329 | -35.6487 | -37.2968
26.639 | -35.6491 | -37.3018
26.6451 | -35.6496 | -37.3079
26.6512 | -35.65 | -37.3127
26.6573 | -35.6504 | -37.3175
26.6634 | -35.6508 | -37.3229
26.6695 | -35.6512 | -37.3279
26.6756 | -35.6516 | -37.3338
26.6817 | -35.6518 | -37.3399
26.6878 | -35.6522 | -37.3455
26.6939 | -35.6527 | -37.3507
26.7 | -35.6529 | -37.3553
26.7061 | -35.6533 | -37.3602
26.7122 | -35.6537 | -37.3656
26.7183 | -35.6541 | -37.3707
26.7245 | -35.6545 | -37.3757
26.7306 | -35.6549 | -37.3817
26.7367 | -35.6553 | -37.3871
26.7428 | -35.6558 | -37.3928
26.7489 | -35.6563 | -37.3986
26.755 | -35.6566 | -37.4043
26.7611 | -35.657 | -37.4087
26.7672 | -35.6573 | -37.4142
26.7733 | -35.6577 | -37.4193
26.7794 | -35.658 | -37.4249
26.7855 | -35.6586 | -37.4301
26.7916 | -35.6589 | -37.4356
26.7977 | -35.6593 | -37.4402
26.8038 | -35.6597 | -37.4449
26.8099 | -35.6601 | -37.4507
26.816 | -35.6604 | -37.4564
26.8221 | -35.6608 | -37.4618
26.8282 | -35.6612 | -37.4668
26.8343 | -35.6617 | -37.4712
26.8404 | -35.662 | -37.4759
26.8465 | -35.6625 | -37.481
26.8526 | -35.6629 | -37.4859
26.8587 | -35.6633 | -37.4903
26.8649 | -35.6638 | -37.4957
26.871 | -35.6642 | -37.5012
26.8771 | -35.6646 | -37.5065
26.8832 | -35.6651 | -37.511
26.8893 | -35.6655 | -37.5165
26.8954 | -35.666 | -37.5212
26.9015 | -35.6665 | -37.5265
26.9076 | -35.6669 | -37.5311
26.9137 | -35.6672 | -37.5366
26.9198 | -35.6677 | -37.5424
26.9259 | -35.6682 | -37.5479
26.932 | -35.6686 | -37.5534
26.9381 | -35.6691 | -37.5589
26.9442 | -35.6696 | -37.5642
26.9503 | -35.67 | -37.5699
26.9564 | -35.6704 | -37.5749
26.9625 | -35.6708 | -37.5797
26.9686 | -35.6712 | -37.5852
26.9747 | -35.6717 | -37.5905
26.9808 | -35.6719 | -37.5957
26.9869 | -35.6724 | -37.6006
26.993 | -35.6729 | -37.6061
26.9991 | -35.6732 | -37.6111
27.0052 | -35.6737 | -37.6154
27.0114 | -35.674 | -37.6199
27.0175 | -35.6743 | -37.6246
27.0236 | -35.6747 | -37.6306
27.0297 | -35.6751 | -37.6353
27.0358 | -35.6757 | -37.6409
27.0419 | -35.6762 | -37.6461
27.048 | -35.6766 | -37.6517
27.0541 | -35.6771 | -37.6574
27.0602 | -35.6775 | -37.6629
27.0663 | -35.6781 | -37.6695
27.0724 | -35.6786 | -37.6739
27.0785 | -35.679 | -37.6791
27.0846 | -35.6793 | -37.6834
27.0907 | -35.6797 | -37.6891
27.0968 | -35.6801 | -37.6945
27.1029 | -35.6804 | -37.699
27.109 | -35.6807 | -37.704
27.1151 | -35.6812 | -37.7101
27.1212 | -35.6817 | -37.7137
27.1273 | -35.682 | -37.7192
27.1334 | -35.6824 | -37.7242
27.1395 | -35.683 | -37.7297
27.1456 | -35.6835 | -37.7351
27.1518 | -35.6839 | -37.741
27.1579 | -35.6843 | -37.746
27.164 | -35.6848 | -37.7509
27.1701 | -35.6853 | -37.7575
27.1762 | -35.6857 | -37.7636
27.1823 | -35.686 | -37.7698
27.1884 | -35.6864 | -37.7748
27.1945 | -35.6867 | -37.778
27.2006 | -35.6869 | -37.7834
27.2067 | -35.6874 | -37.788
27.2128 | -35.6876 | -37.7933
27.2189 | -35.688 | -37.7982
27.225 | -35.6885 | -37.8038
27.2311 | -35.6889 | -37.8078
27.2372 | -35.6894 | -37.8129
27.2433 | -35.6898 | -37.818
27.2494 | -35.6901 | -37.8238
27.2555 | -35.6906 | -37.8293
27.2616 | -35.691 | -37.8328
27.2677 | -35.6914 | -37.8376
27.2738 | -35.6919 | -37.8427
27.2799 | -35.6924 | -37.8485
27.286 | -35.6927 | -37.8528
27.2921 | -35.6931 | -37.8584
27.2983 | -35.6933 | -37.8637
27.3044 | -35.6937 | -37.8682
27.3105 | -35.6942 | -37.8733
27.3166 | -35.6946 | -37.8783
27.3227 | -35.6949 | -37.8836
27.3288 | -35.6954 | -37.8891
27.3349 | -35.6957 | -37.894
27.341 | -35.6962 | -37.899
27.3471 | -35.6966 | -37.9041
27.3532 | -35.6969 | -37.9095
27.3593 | -35.6974 | -37.9151
27.3654 | -35.6979 | -37.9201
27.3715 | -35.6984 | -37.9247
27.3776 | -35.6987 | -37.9299
27.3837 | -35.6991 | -37.9351
27.3898 | -35.6996 | -37.9391
27.3959 | -35.7 | -37.9434
27.402 | -35.7003 | -37.9467
27.4081 | -35.7007 | -37.9506
27.4142 | -35.7012 | -37.9556
27.4203 | -35.7015 | -37.9607
27.4264 | -35.7019 | -37.9664
27.4325 | -35.7022 | -37.9707
27.4387 | -35.7027 | -37.9756
27.4448 | -35.7031 | -37.9812
27.4509 | -35.7035 | -37.986
27.457 | -35.7039 | -37.9916
27.4631 | -35.7044 | -37.9967
27.4692 | -35.7046 | -38.0017
27.4753 | -35.7051 | -38.0071
27.4814 | -35.7054 | -38.0122
27.4875 | -35.7057 | -38.0174
27.4936 | -35.7061 | -38.0228
27.4997 | -35.7065 | -38.0284
27.5058 | -35.7069 | -38.0334
27.5119 | -35.7073 | -38.0385
27.518 | -35.7077 | -38.044
27.5241 | -35.708 | -38.0496
27.5302 | -35.7084 | -38.0545
27.5363 | -35.7087 | -38.0598
27.5424 | -35.709 | -38.0644
27.5485 | -35.7095 | -38.0694
27.5546 | -35.7098 | -38.0741
27.5607 | -35.7102 | -38.0792
27.5668 | -35.7105 | -38.0847
27.5729 | -35.7109 | -38.0901
27.5791 | -35.7113 | -38.0949
27.5852 | -35.7118 | -38.0994
27.5913 | -35.7122 | -38.1048
27.5974 | -35.7125 | -38.1103
27.6035 | -35.7128 | -38.1157
27.6096 | -35.7131 | -38.1214
27.6157 | -35.7135 | -38.1261
27.6218 | -35.7139 | -38.1314
27.6279 | -35.7143 | -38.1367
27.634 | -35.7147 | -38.142
27.6401 | -35.7151 | -38.1473
27.6462 | -35.7155 | -38.1514
27.6523 | -35.7158 | -38.1564
27.6584 | -35.7162 | -38.1611
27.6645 | -35.7166 | -38.166
27.6706 | -35.7169 | -38.1707
27.6767 | -35.7173 | -38.1755
27.6828 | -35.7178 | -38.1801
27.6889 | -35.7181 | -38.1848
27.695 | -35.7184 | -38.1901
27.7011 | -35.7188 | -38.1955
27.7072 | -35.7191 | -38.2001
27.7133 | -35.7194 | -38.205
27.7194 | -35.7197 | -38.2097
27.7256 | -35.72 | -38.2137
27.7317 | -35.7203 | -38.2184
27.7378 | -35.7207 | -38.2235
27.7439 | -35.7213 | -38.2287
27.75 | -35.7216 | -38.2342
27.7561 | -35.7219 | -38.2399
27.7622 | -35.7223 | -38.2448
27.7683 | -35.7226 | -38.2496
27.7744 | -35.7232 | -38.2553
27.7805 | -35.7234 | -38.2606
27.7866 | -35.7236 | -38.2656
27.7927 | -35.7239 | -38.2704
27.7988 | -35.7244 | -38.2762
27.8049 | -35.7247 | -38.2828
27.811 | -35.7251 | -38.2878
27.8171 | -35.7258 | -38.2932
27.8232 | -35.726 | -38.2987
27.8293 | -35.7263 | -38.3037
27.8354 | -35.7265 | -38.308
27.8415 | -35.7269 | -38.314
27.8476 | -35.7273 | -38.3184
27.8537 | -35.7277 | -38.3244
27.8598 | -35.728 | -38.329
27.866 | -35.7284 | -38.3337
27.8721 | -35.7287 | -38.3385
27.8782 | -35.7291 | -38.3431
27.8843 | -35.7295 | -38.3483
27.8904 | -35.7299 | -38.3534
27.8965 | -35.7302 | -38.3582
27.9026 | -35.7306 | -38.3633
27.9087 | -35.731 | -38.3682
27.9148 | -35.7314 | -38.3726
27.9209 | -35.7318 | -38.3779
27.927 | -35.7322 | -38.384
27.9331 | -35.7324 | -38.3893
27.9392 | -35.7325 | -38.3948
27.9453 | -35.7329 | -38.3989
27.9514 | -35.7332 | -38.4039
27.9575 | -35.7336 | -38.4091
27.9636 | -35.7338 | -38.4138
27.9697 | -35.7341 | -38.4215
27.9758 | -35.7346 | -38.4273
27.9819 | -35.7351 | -38.4323
27.988 | -35.7354 | -38.4366
27.9941 | -35.7358 | -38.4417
28.0002 | -35.7361 | -38.4471
28.0063 | -35.7364 | -38.4518
28.0125 | -35.7367 | -38.4567
28.0186 | -35.7369 | -38.4615
28.0247 | -35.7373 | -38.4666
28.0308 | -35.7376 | -38.4712
28.0369 | -35.7379 | -38.4757
28.043 | -35.7382 | -38.481
28.0491 | -35.7387 | -38.4876
28.0552 | -35.739 | -38.4932
28.0613 | -35.7394 | -38.4978
28.0674 | -35.7398 | -38.5022
28.0735 | -35.7402 | -38.5068

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5327 | -43.3895
21.4908 | -35.5404 | -35.4455
16.4948 | -34.9198 | -27.9427
12.6602 | -34.5586 | -21.5905
9.7171 | -34.3434 | -16.6354
7.4582 | -34.1933 | -12.6237
5.7244 | -34.0905 | -9.3733
4.3936 | -34.0286 | -6.762
3.3722 | -34.017 | -4.6986
2.5883 | -34.001 | -3.1329
1.9866 | -33.9901 | -1.883
1.5248 | -33.9997 | -0.88601
1.1703 | -34.0101 | -0.22981
1 | -33.9886 | 0.091689
0.89822 | -34.0046 | 0.3652
0.68941 | -33.991 | 0.95014
0.52919 | -33.9793 | 1.3998
0.40612 | -33.975 | 2.0099
0.31171 | -33.9913 | 2.4593
0.23928 | -33.9967 | 3.189
0.18358 | -34.0187 | 3.6083
0.14091 | -34.0321 | 4.4959
0.10819 | -34.0678 | 5.6507
0.083017 | -34.1066 | 7.1232
0.063705 | -34.1474 | 8.9947
0.048923 | -34.2238 | 11.4807
0.037479 | -34.3273 | 14.6774
0.028801 | -34.5164 | 18.6784
0.022078 | -34.8237 | 23.9527
0.017023 | -35.2963 | 30.3048
0.013018 | -36.0796 | 37.9364
0.010014 | -37.1836 | 45.8306

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003689 | -8.2534
0.007377 | -26.4646
0.011066 | -35.5345
0.014754 | -40.2852
0.018443 | -43.1871
0.022131 | -43.7348
0.02582 | -45.6387
0.029508 | -47.8842
0.033197 | -47.6208
0.036885 | -48.3564
0.040574 | -49.8444
0.044262 | -51.7527
0.047951 | -53.4238
0.051639 | -53.3994
0.055328 | -54.0444
0.059016 | -56.2144
0.062705 | -56.7681
0.066393 | -57.2882
0.070082 | -57.6494
0.07377 | -58.4866
0.077459 | -59.7872
0.081147 | -60.0189
0.084836 | -60.193
0.088524 | -60.6285
0.092213 | -62.2152
0.095901 | -62.6767
0.09959 | -61.8209
0.10328 | -62.7532
0.10697 | -62.4772
0.11066 | -62.2415
0.11434 | -63.4465
0.11803 | -65.1288
0.12172 | -65.998
0.12541 | -66.0891
0.1291 | -65.2798
0.13279 | -65.5135
0.13648 | -65.8111
0.14016 | -65.7682
0.14385 | -65.8567
0.14754 | -66.7613
0.15123 | -67.587
0.15492 | -66.5799
0.15861 | -66.1765
0.16229 | -67.0584
0.16598 | -68.2506
0.16967 | -67.8965
0.17336 | -68.5736
0.17705 | -69.5201
0.18074 | -69.0212
0.18443 | -69.5262
0.18811 | -69.8512
0.1918 | -70.2553
0.19549 | -71.438
0.19918 | -70.7537
0.20287 | -71.168
0.20656 | -70.8322
0.21024 | -70.5447
0.21393 | -70.4186
0.21762 | -71.6508
0.22131 | -71.2468
0.225 | -71.015
0.22869 | -71.6017
0.23238 | -72.9277
0.23606 | -72.3363
0.23975 | -72.423
0.24344 | -72.9484
0.24713 | -73.4451
0.25082 | -73.4792
0.25451 | -73.7824
0.2582 | -74.2061
0.26188 | -73.616
0.26557 | -72.8988
0.26926 | -72.745
0.27295 | -73.5787
0.27664 | -75.0291
0.28033 | -75.632
0.28402 | -75.1391
0.2877 | -74.9023
0.29139 | -74.4784
0.29508 | -74.5936
0.29877 | -74.5626
0.30246 | -76.1693
0.30615 | -75.6843
0.30983 | -76.3487
0.31352 | -75.5961
0.31721 | -75.8031
0.3209 | -76.0015
0.32459 | -75.0463
0.32828 | -75.9224
0.33197 | -75.7341
0.33565 | -75.399
0.33934 | -75.905
0.34303 | -75.9308
0.34672 | -76.5737
0.35041 | -77.3338
0.3541 | -77.8561
0.35779 | -77.6843
0.36147 | -77.4653
0.36516 | -76.8401
0.36885 | -77.1502
0.37254 | -77.0667
0.37623 | -77.0105
0.37992 | -76.3485
0.3836 | -76.3832
0.38729 | -77.3991
0.39098 | -77.5375
0.39467 | -78.7212
0.39836 | -78.5591
0.40205 | -77.9857
0.40574 | -78.04
0.40942 | -77.9766
0.41311 | -77.9616
0.4168 | -78.0321
0.42049 | -77.683
0.42418 | -77.1869
0.42787 | -78.1835
0.43156 | -79.0348
0.43524 | -80.1779
0.43893 | -79.8267
0.44262 | -78.5426
0.44631 | -79.0339
0.45 | -79.6388
0.45369 | -79.1369
0.45737 | -79.3632
0.46106 | -79.7899
0.46475 | -79.6961
0.46844 | -78.2711
0.47213 | -78.5761
0.47582 | -78.9465
0.47951 | -79.8273
0.48319 | -79.2642
0.48688 | -79.1009
0.49057 | -80.2157
0.49426 | -80.6451
0.49795 | -79.9465
0.50164 | -81.0513
0.50533 | -81.0193
0.50901 | -80.9017
0.5127 | -79.6538
0.51639 | -79.3878
0.52008 | -80.3943
0.52377 | -80.505
0.52746 | -80.0987
0.53114 | -80.8143
0.53483 | -81.2429
0.53852 | -81.3807
0.54221 | -81.138
0.5459 | -80.9051
0.54959 | -80.9022
0.55328 | -80.6294
0.55696 | -80.6031
0.56065 | -81.0762
0.56434 | -81.8443
0.56803 | -82.3469
0.57172 | -82.2704
0.57541 | -81.8516
0.5791 | -81.5294
0.58278 | -81.6412
0.58647 | -82.1351
0.59016 | -82.2066
0.59385 | -82.0258
0.59754 | -82.1282
0.60123 | -82.428
0.60491 | -82.565
0.6086 | -82.4224
0.61229 | -82.6334
0.61598 | -82.8594
0.61967 | -82.6203
0.62336 | -82.6262
0.62705 | -83.0047
0.63073 | -82.9192
0.63442 | -82.594
0.63811 | -82.4335
0.6418 | -82.357
0.64549 | -82.4085
0.64918 | -82.416
0.65287 | -82.4494
0.65655 | -82.6022
0.66024 | -82.8061
0.66393 | -82.8312
0.66762 | -82.5118
0.67131 | -82.3823
0.675 | -82.4514
0.67868 | -82.8423
0.68237 | -83.0365
0.68606 | -82.7516
0.68975 | -82.8638
0.69344 | -83.3788
0.69713 | -83.4987
0.70082 | -83.5505
0.70451 | -83.8604
0.70819 | -84.2443
0.71188 | -84.5375
0.71557 | -84.6118
0.71926 | -84.4403
0.72295 | -84.5202
0.72664 | -84.6119
0.73032 | -84.1343
0.73401 | -83.2891
0.7377 | -82.8742
0.74139 | -83.3621
0.74508 | -84.3678
0.74877 | -84.4999
0.75245 | -84.4176
0.75614 | -84.4815
0.75983 | -84.6522
0.76352 | -85.0482
0.76721 | -85.2759
0.7709 | -84.9368
0.77459 | -84.5106
0.77828 | -84.4869
0.78196 | -84.2893
0.78565 | -83.6599
0.78934 | -83.5267
0.79303 | -84.1074
0.79672 | -84.7584
0.80041 | -85.0878
0.80409 | -85.2666
0.80778 | -85.4347
0.81147 | -85.6096
0.81516 | -85.6232
0.81885 | -85.6185
0.82254 | -85.6122
0.82622 | -85.4864
0.82991 | -85.5038
0.8336 | -85.4181
0.83729 | -84.9588
0.84098 | -84.9238
0.84467 | -85.1822
0.84836 | -85.0189
0.85205 | -85.2417
0.85573 | -85.6779
0.85942 | -85.5198
0.86311 | -85.3682
0.8668 | -85.4767
0.87049 | -85.7631
0.87418 | -86.28
0.87786 | -86.3157
0.88155 | -86.4582
0.88524 | -86.8343
0.88893 | -86.816
0.89262 | -86.5331
0.89631 | -86.3659
0.89999 | -86.1156
0.90368 | -86.1367
0.90737 | -86.3099
0.91106 | -86.2467
0.91475 | -86.101
0.91844 | -85.7322
0.92213 | -85.6731
0.92582 | -86.2234
0.9295 | -86.7503
0.93319 | -86.8429
0.93688 | -86.857
0.94057 | -86.8535
0.94426 | -86.704
0.94795 | -86.7711
0.95163 | -86.8701
0.95532 | -86.702
0.95901 | -86.507
0.9627 | -86.2356
0.96639 | -85.7758
0.97008 | -85.4601
0.97377 | -85.6174
0.97745 | -86.3175
0.98114 | -87.4323
0.98483 | -88.2407
0.98852 | -88.0132
0.99221 | -87.6798
0.9959 | -87.77
0.99958 | -88.0796
1.0033 | -88.2147
1.007 | -87.9209
1.0107 | -87.1184
1.0143 | -86.393
1.018 | -86.0534
1.0217 | -86.477
1.0254 | -87.4276
1.0291 | -87.6526
1.0328 | -87.5135
1.0365 | -87.5434
1.0402 | -87.2822
1.0438 | -87.1163
1.0475 | -87.1359
1.0512 | -87.3873
1.0549 | -88.0766
1.0586 | -88.1148
1.0623 | -87.7994
1.066 | -87.9837
1.0697 | -88.4111
1.0734 | -88.5442
1.077 | -88.2241
1.0807 | -87.7834
1.0844 | -87.4927
1.0881 | -87.5122
1.0918 | -87.8768
1.0955 | -88.3212
1.0992 | -88.5549
1.1029 | -88.595
1.1066 | -88.4978
1.1102 | -88.3289
1.1139 | -88.2426
1.1176 | -88.3214
1.1213 | -88.5059
1.125 | -88.6724
1.1287 | -88.7429
1.1324 | -88.5702
1.1361 | -88.1592
1.1397 | -87.8014
1.1434 | -87.6251
1.1471 | -87.4812
1.1508 | -87.2871
1.1545 | -87.2623
1.1582 | -87.5732
1.1619 | -88.049
1.1656 | -88.4531
1.1693 | -88.7457
1.1729 | -88.9756
1.1766 | -89.1747
1.1803 | -89.3048
1.184 | -89.2271
1.1877 | -89.043
1.1914 | -88.9312
1.1951 | -88.8212
1.1988 | -88.7397
1.2025 | -88.8457
1.2061 | -89.0119
1.2098 | -88.929
1.2135 | -88.6315
1.2172 | -88.468
1.2209 | -88.4475
1.2246 | -88.3577
1.2283 | -88.2556
1.232 | -88.2708
1.2356 | -88.3722
1.2393 | -88.6152
1.243 | -89.0673
1.2467 | -89.5968
1.2504 | -90.0362
1.2541 | -90.3072
1.2578 | -90.2996
1.2615 | -90.0461
1.2652 | -89.7496
1.2688 | -89.5139
1.2725 | -89.3276
1.2762 | -89.2116
1.2799 | -89.1178
1.2836 | -88.9812
1.2873 | -88.8727
1.291 | -88.8662
1.2947 | -88.9388
1.2984 | -89.064
1.302 | -89.2624
1.3057 | -89.5351
1.3094 | -89.8297
1.3131 | -89.9872
1.3168 | -89.7126
1.3205 | -89.1704
1.3242 | -88.9189
1.3279 | -89.0558
1.3316 | -89.2237
1.3352 | -89.1551
1.3389 | -89.01
1.3426 | -89.0285
1.3463 | -89.2082
1.35 | -89.4396
1.3537 | -89.7134
1.3574 | -90.0274
1.3611 | -90.2091
1.3647 | -90.1251
1.3684 | -89.953
1.3721 | -89.7786
1.3758 | -89.5266
1.3795 | -89.2999
1.3832 | -89.2815
1.3869 | -89.5151
1.3906 | -89.9865
1.3943 | -90.5522
1.3979 | -90.8369
1.4016 | -90.7432
1.4053 | -90.5618
1.409 | -90.3224
1.4127 | -89.9712
1.4164 | -89.762
1.4201 | -89.8661
1.4238 | -90.0915
1.4275 | -90.1752
1.4311 | -90.145
1.4348 | -90.1674
1.4385 | -90.2553
1.4422 | -90.3597
1.4459 | -90.5042
1.4496 | -90.6272
1.4533 | -90.575
1.457 | -90.3819
1.4606 | -90.2816
1.4643 | -90.4176
1.468 | -90.6595
1.4717 | -90.7235
1.4754 | -90.457
1.4791 | -90.0168
1.4828 | -89.7926
1.4865 | -89.9727
1.4902 | -90.3292
1.4938 | -90.5197
1.4975 | -90.4533
1.5012 | -90.3071
1.5049 | -90.2985
1.5086 | -90.4916
1.5123 | -90.826
1.516 | -91.0964
1.5197 | -91.0388
1.5234 | -90.7256
1.527 | -90.4935
1.5307 | -90.5329
1.5344 | -90.7636
1.5381 | -90.8686
1.5418 | -90.7007
1.5455 | -90.5402
1.5492 | -90.6008
1.5529 | -90.7609
1.5565 | -90.7957
1.5602 | -90.7502
1.5639 | -90.8513
1.5676 | -91.025
1.5713 | -91.1191
1.575 | -91.2749
1.5787 | -91.48
1.5824 | -91.5406
1.5861 | -91.5262
1.5897 | -91.5394
1.5934 | -91.424
1.5971 | -91.3324
1.6008 | -91.3024
1.6045 | -91.3731
1.6082 | -91.496
1.6119 | -91.6134
1.6156 | -91.7076
1.6193 | -91.7663
1.6229 | -91.7905
1.6266 | -91.7932
1.6303 | -91.7826
1.634 | -91.6986
1.6377 | -91.5248
1.6414 | -91.3737
1.6451 | -91.3292
1.6488 | -91.3727
1.6525 | -91.4178
1.6561 | -91.3689
1.6598 | -91.2645
1.6635 | -91.2448
1.6672 | -91.4296
1.6709 | -91.7764
1.6746 | -92.1262
1.6783 | -92.3479
1.682 | -92.3592
1.6856 | -92.1881
1.6893 | -91.9352
1.693 | -91.6738
1.6967 | -91.4324
1.7004 | -91.2404
1.7041 | -91.1662
1.7078 | -91.2362
1.7115 | -91.3997
1.7152 | -91.5836
1.7188 | -91.7709
1.7225 | -92.0285
1.7262 | -92.3698
1.7299 | -92.6935
1.7336 | -92.8322
1.7373 | -92.7051
1.741 | -92.4267
1.7447 | -92.1928
1.7484 | -92.0759
1.752 | -92.0078
1.7557 | -91.9253
1.7594 | -91.8734
1.7631 | -91.9232
1.7668 | -92.0622
1.7705 | -92.2052
1.7742 | -92.2579
1.7779 | -92.1667
1.7815 | -91.9754
1.7852 | -91.7999
1.7889 | -91.7491
1.7926 | -91.858
1.7963 | -92.0519
1.8 | -92.2433
1.8037 | -92.4132
1.8074 | -92.5799
1.8111 | -92.7265
1.8147 | -92.8164
1.8184 | -92.8053
1.8221 | -92.7188
1.8258 | -92.6742
1.8295 | -92.724
1.8332 | -92.7787
1.8369 | -92.7388
1.8406 | -92.6134
1.8443 | -92.4573
1.8479 | -92.3264
1.8516 | -92.3022
1.8553 | -92.4097
1.859 | -92.5935
1.8627 | -92.7874
1.8664 | -92.9342
1.8701 | -93.0147
1.8738 | -93.0334
1.8774 | -92.9887
1.8811 | -92.8685
1.8848 | -92.6872
1.8885 | -92.5379
1.8922 | -92.4634
1.8959 | -92.442
1.8996 | -92.4079
1.9033 | -92.3089
1.907 | -92.1743
1.9106 | -92.0931
1.9143 | -92.1589
1.918 | -92.409
1.9217 | -92.7945
1.9254 | -93.2144
1.9291 | -93.5226
1.9328 | -93.5994
1.9365 | -93.4056
1.9402 | -92.9945
1.9438 | -92.5604
1.9475 | -92.343
1.9512 | -92.4416
1.9549 | -92.7383
1.9586 | -93.0459
1.9623 | -93.2776
1.966 | -93.4545
1.9697 | -93.5878
1.9734 | -93.6283
1.977 | -93.5473
1.9807 | -93.3646
1.9844 | -93.119
1.9881 | -92.9119
1.9918 | -92.8244
1.9955 | -92.8481
1.9992 | -92.9124
2.0029 | -92.9613
2.0065 | -93.0458
2.0102 | -93.2227
2.0139 | -93.4809
2.0176 | -93.7498
2.0213 | -93.923
2.025 | -94.0032
2.0287 | -94.0762
2.0324 | -94.1324
2.0361 | -94.1125
2.0397 | -93.9882
2.0434 | -93.8204
2.0471 | -93.6871
2.0508 | -93.6119
2.0545 | -93.5921
2.0582 | -93.5925
2.0619 | -93.5473
2.0656 | -93.4732
2.0693 | -93.4572
2.0729 | -93.5384
2.0766 | -93.6711
2.0803 | -93.8095
2.084 | -93.9821
2.0877 | -94.1777
2.0914 | -94.351
2.0951 | -94.4994
2.0988 | -94.6501
2.1024 | -94.7443
2.1061 | -94.6685
2.1098 | -94.4072
2.1135 | -94.0786
2.1172 | -93.8172
2.1209 | -93.7069
2.1246 | -93.7418
2.1283 | -93.8626
2.132 | -93.9302
2.1356 | -93.9532
2.1393 | -93.9132
2.143 | -93.841
2.1467 | -93.7811
2.1504 | -93.7659
2.1541 | -93.7958
2.1578 | -93.83
2.1615 | -93.8385
2.1652 | -93.8266
2.1688 | -93.7933
2.1725 | -93.7322
2.1762 | -93.6373
2.1799 | -93.5351
2.1836 | -93.4374
2.1873 | -93.351
2.191 | -93.2868
2.1947 | -93.2639
2.1983 | -93.3247
2.202 | -93.5048
2.2057 | -93.7841
2.2094 | -94.1126
2.2131 | -94.3992
2.2168 | -94.5671
2.2205 | -94.5558
2.2242 | -94.4159
2.2279 | -94.2639
2.2315 | -94.1996
2.2352 | -94.2588
2.2389 | -94.4084
2.2426 | -94.5543
2.2463 | -94.6329
2.25 | -94.6288
2.2537 | -94.5667
2.2574 | -94.4816
2.2611 | -94.3956
2.2647 | -94.3002
2.2684 | -94.196
2.2721 | -94.1221
2.2758 | -94.0998
2.2795 | -94.1159
2.2832 | -94.1653
2.2869 | -94.2405
2.2906 | -94.346
2.2943 | -94.4858
2.2979 | -94.6327
2.3016 | -94.7286
2.3053 | -94.7227
2.309 | -94.6224
2.3127 | -94.4736
2.3164 | -94.3203
2.3201 | -94.194
2.3238 | -94.1032
2.3274 | -94.0589
2.3311 | -94.0477
2.3348 | -94.0196
2.3385 | -93.9349
2.3422 | -93.7949
2.3459 | -93.6532
2.3496 | -93.5701
2.3533 | -93.5712
2.357 | -93.6213
2.3606 | -93.6522
2.3643 | -93.6302
2.368 | -93.5731
2.3717 | -93.5578
2.3754 | -93.6737
2.3791 | -93.9344
2.3828 | -94.2849
2.3865 | -94.6162
2.3902 | -94.8652
2.3938 | -95.0264
2.3975 | -95.1159
2.4012 | -95.1152
2.4049 | -95.0257
2.4086 | -94.8988
2.4123 | -94.7975
2.416 | -94.7519
2.4197 | -94.7644
2.4233 | -94.8188
2.427 | -94.8988
2.4307 | -94.9891
2.4344 | -95.0812
2.4381 | -95.1458
2.4418 | -95.1671
2.4455 | -95.1514
2.4492 | -95.1389
2.4529 | -95.1719
2.4565 | -95.2576
2.4602 | -95.3686
2.4639 | -95.4557
2.4676 | -95.4821
2.4713 | -95.447
2.475 | -95.3966
2.4787 | -95.3817
2.4824 | -95.4019
2.4861 | -95.4508
2.4897 | -95.5106
2.4934 | -95.5547
2.4971 | -95.5341
2.5008 | -95.4141
2.5045 | -95.2431
2.5082 | -95.1439
2.5119 | -95.1894
2.5156 | -95.3486
2.5192 | -95.5333
2.5229 | -95.6565
2.5266 | -95.6602
2.5303 | -95.5758
2.534 | -95.4735
2.5377 | -95.384
2.5414 | -95.3088
2.5451 | -95.2321
2.5488 | -95.1717
2.5524 | -95.1584
2.5561 | -95.1862
2.5598 | -95.2335
2.5635 | -95.2883
2.5672 | -95.3732
2.5709 | -95.4954
2.5746 | -95.6372
2.5783 | -95.7732
2.582 | -95.8764
2.5856 | -95.9403
2.5893 | -95.972
2.593 | -95.9939
2.5967 | -96.0237
2.6004 | -96.0609
2.6041 | -96.0442
2.6078 | -95.9149
2.6115 | -95.684
2.6152 | -95.4313
2.6188 | -95.2319
2.6225 | -95.1204
2.6262 | -95.0988
2.6299 | -95.1492
2.6336 | -95.2602
2.6373 | -95.3928
2.641 | -95.4937
2.6447 | -95.5457
2.6483 | -95.5735
2.652 | -95.5943
2.6557 | -95.6024
2.6594 | -95.5334
2.6631 | -95.4788
2.6668 | -95.4177
2.6705 | -95.3914
2.6742 | -95.4259
2.6779 | -95.5252
2.6815 | -95.6771
2.6852 | -95.8518
2.6889 | -96.0143
2.6926 | -96.104
2.6963 | -96.1163
2.7 | -96.0589
2.7037 | -95.9593
2.7074 | -95.8543
2.7111 | -95.7915
2.7147 | -95.7896
2.7184 | -95.8317
2.7221 | -95.9136
2.7258 | -96.0078
2.7295 | -96.0958
2.7332 | -96.1478
2.7369 | -96.1502
2.7406 | -96.0972
2.7442 | -95.9921
2.7479 | -95.8704
2.7516 | -95.7709
2.7553 | -95.7094
2.759 | -95.684
2.7627 | -95.6966
2.7664 | -95.7231
2.7701 | -95.7269
2.7738 | -95.7099
2.7774 | -95.6868
2.7811 | -95.7046
2.7848 | -95.7845
2.7885 | -95.908
2.7922 | -96.0407
2.7959 | -96.1371
2.7996 | -96.1712
2.8033 | -96.1637
2.807 | -96.1346
2.8106 | -96.0991
2.8143 | -96.0838
2.818 | -96.1092
2.8217 | -96.1743
2.8254 | -96.2673
2.8291 | -96.3589
2.8328 | -96.4111
2.8365 | -96.3935
2.8401 | -96.2762
2.8438 | -96.0835
2.8475 | -95.8603
2.8512 | -95.6474
2.8549 | -95.4866
2.8586 | -95.4124
2.8623 | -95.4322
2.866 | -95.5273
2.8697 | -95.6822
2.8733 | -95.8624
2.877 | -96.0392
2.8807 | -96.2118
2.8844 | -96.3794
2.8881 | -96.543
2.8918 | -96.6731
2.8955 | -96.7614
2.8992 | -96.8101
2.9029 | -96.8298
2.9065 | -96.8328
2.9102 | -96.8251
2.9139 | -96.8031
2.9176 | -96.7621
2.9213 | -96.7243
2.925 | -96.7034
2.9287 | -96.7041
2.9324 | -96.7251
2.9361 | -96.7585
2.9397 | -96.8129
2.9434 | -96.8684
2.9471 | -96.9063
2.9508 | -96.9074
2.9545 | -96.8646
2.9582 | -96.7942
2.9619 | -96.7126
2.9656 | -96.6506
2.9692 | -96.6003
2.9729 | -96.5541
2.9766 | -96.4961
2.9803 | -96.4029
2.984 | -96.2796
2.9877 | -96.152
2.9914 | -96.0459
2.9951 | -95.9991
2.9988 | -96.0225
3.0024 | -96.107
3.0061 | -96.2408
3.0098 | -96.3793
3.0135 | -96.4725
3.0172 | -96.5024
3.0209 | -96.4682
3.0246 | -96.4024
3.0283 | -96.3425
3.032 | -96.311
3.0356 | -96.3101
3.0393 | -96.3379
3.043 | -96.3636
3.0467 | -96.3872
3.0504 | -96.4125
3.0541 | -96.4457
3.0578 | -96.4949
3.0615 | -96.5394
3.0651 | -96.5549
3.0688 | -96.5152
3.0725 | -96.4269
3.0762 | -96.3348
3.0799 | -96.2726
3.0836 | -96.2566
3.0873 | -96.2802
3.091 | -96.3094
3.0947 | -96.3194
3.0983 | -96.3146
3.102 | -96.305
3.1057 | -96.314
3.1094 | -96.3563
3.1131 | -96.4287
3.1168 | -96.507
3.1205 | -96.5705
3.1242 | -96.6119
3.1279 | -96.6349
3.1315 | -96.6535
3.1352 | -96.6776
3.1389 | -96.7126
3.1426 | -96.7636
3.1463 | -96.8267
3.15 | -96.9026
3.1537 | -96.9833
3.1574 | -97.0651
3.161 | -97.1523
3.1647 | -97.2483
3.1684 | -97.3358
3.1721 | -97.3996
3.1758 | -97.4323
3.1795 | -97.4369
3.1832 | -97.3879
3.1869 | -97.284
3.1906 | -97.0888
3.1942 | -96.9405
3.1979 | -96.8068
3.2016 | -96.7105
3.2053 | -96.6707
3.209 | -96.6782
3.2127 | -96.7141
3.2164 | -96.7692
3.2201 | -96.8094
3.2238 | -96.8287
3.2274 | -96.8315
3.2311 | -96.8059
3.2348 | -96.7724
3.2385 | -96.7322
3.2422 | -96.6961
3.2459 | -96.6771
3.2496 | -96.6801
3.2533 | -96.718
3.257 | -96.7825
3.2606 | -96.8483
3.2643 | -96.8959
3.268 | -96.9195
3.2717 | -96.9118
3.2754 | -96.889
3.2791 | -96.8568
3.2828 | -96.8335
3.2865 | -96.8489
3.2901 | -96.9112
3.2938 | -97.0154
3.2975 | -97.1469
3.3012 | -97.2757
3.3049 | -97.3931
3.3086 | -97.4744
3.3123 | -97.5054
3.316 | -97.4984
3.3197 | -97.4617
3.3233 | -97.3905
3.327 | -97.3092
3.3307 | -97.2308
3.3344 | -97.1664
3.3381 | -97.1251
3.3418 | -97.097
3.3455 | -97.0776
3.3492 | -97.0714
3.3529 | -97.0716
3.3565 | -97.0765
3.3602 | -97.0838
3.3639 | -97.1065
3.3676 | -97.1479
3.3713 | -97.2117
3.375 | -97.2919
3.3787 | -97.3805
3.3824 | -97.4613
3.386 | -97.5129
3.3897 | -97.5216
3.3934 | -97.4854
3.3971 | -97.4219
3.4008 | -97.3542
3.4045 | -97.2893
3.4082 | -97.2478
3.4119 | -97.2362
3.4156 | -97.2419
3.4192 | -97.2712
3.4229 | -97.3277
3.4266 | -97.4014
3.4303 | -97.5072
3.434 | -97.6406
3.4377 | -97.764
3.4414 | -97.8606
3.4451 | -97.924
3.4488 | -97.9494
3.4524 | -97.9405
3.4561 | -97.8963
3.4598 | -97.8434
3.4635 | -97.7849
3.4672 | -97.7366
3.4709 | -97.7262
3.4746 | -97.7373
3.4783 | -97.7551
3.4819 | -97.7573
3.4856 | -97.738
3.4893 | -97.7042
3.493 | -97.6907
3.4967 | -97.7035
3.5004 | -97.7296
3.5041 | -97.7793
3.5078 | -97.8299
3.5115 | -97.8816
3.5151 | -97.9096
3.5188 | -97.9179
3.5225 | -97.9056
3.5262 | -97.8737
3.5299 | -97.8306
3.5336 | -97.78
3.5373 | -97.7225
3.541 | -97.6585
3.5447 | -97.6004
3.5483 | -97.5428
3.552 | -97.4941
3.5557 | -97.4619
3.5594 | -97.4413
3.5631 | -97.4275
3.5668 | -97.4076
3.5705 | -97.3925
3.5742 | -97.3864
3.5779 | -97.3936
3.5815 | -97.4151
3.5852 | -97.4541
3.5889 | -97.5049
3.5926 | -97.5688
3.5963 | -97.6343
3.6 | -97.7042
3.6037 | -97.7891
3.6074 | -97.8814
3.611 | -97.9742
3.6147 | -98.0567
3.6184 | -98.1286
3.6221 | -98.1884
3.6258 | -98.2235
3.6295 | -98.2221
3.6332 | -98.1714
3.6369 | -98.0689
3.6406 | -97.9314
3.6442 | -97.7872
3.6479 | -97.6615
3.6516 | -97.5859
3.6553 | -97.5595
3.659 | -97.5843
3.6627 | -97.6588
3.6664 | -97.7717
3.6701 | -97.9201
3.6738 | -98.0668
3.6774 | -98.1732
3.6811 | -98.2304
3.6848 | -98.2265
3.6885 | -98.179
3.6922 | -98.1172
3.6959 | -98.0621
3.6996 | -98.0289
3.7033 | -98.0314
3.7069 | -98.0608
3.7106 | -98.1177
3.7143 | -98.1877
3.718 | -98.2493
3.7217 | -98.2891
3.7254 | -98.3324
3.7291 | -98.3632
3.7328 | -98.392
3.7365 | -98.4137
3.7401 | -98.432
3.7438 | -98.4444
3.7475 | -98.4595
3.7512 | -98.4792
3.7549 | -98.4961
3.7586 | -98.5132
3.7623 | -98.5279
3.766 | -98.538
3.7697 | -98.5289
3.7733 | -98.5039
3.777 | -98.4713
3.7807 | -98.425
3.7844 | -98.3808
3.7881 | -98.3558
3.7918 | -98.3445
3.7955 | -98.3467
3.7992 | -98.3635
3.8028 | -98.3749
3.8065 | -98.3938
3.8102 | -98.4109
3.8139 | -98.4286
3.8176 | -98.4596
3.8213 | -98.4899
3.825 | -98.5121
3.8287 | -98.5194
3.8324 | -98.5193
3.836 | -98.5023
3.8397 | -98.4724
3.8434 | -98.4363
3.8471 | -98.3969
3.8508 | -98.3683
3.8545 | -98.3568
3.8582 | -98.3674
3.8619 | -98.3995
3.8656 | -98.4502
3.8692 | -98.5089
3.8729 | -98.5758
3.8766 | -98.6265
3.8803 | -98.6627
3.884 | -98.6803
3.8877 | -98.6812
3.8914 | -98.6695
3.8951 | -98.6463
3.8988 | -98.6285
3.9024 | -98.6222
3.9061 | -98.6352
3.9098 | -98.6702
3.9135 | -98.73
3.9172 | -98.7983
3.9209 | -98.8647
3.9246 | -98.9132
3.9283 | -98.9263
3.9319 | -98.8981
3.9356 | -98.8333
3.9393 | -98.7613
3.943 | -98.6885
3.9467 | -98.6389
3.9504 | -98.6173
3.9541 | -98.6257
3.9578 | -98.6592
3.9615 | -98.6956
3.9651 | -98.7252
3.9688 | -98.7499
3.9725 | -98.7692
3.9762 | -98.7878
3.9799 | -98.8303
3.9836 | -98.8899
3.9873 | -98.968
3.991 | -99.0449
3.9947 | -99.1049
3.9983 | -99.1418
4.002 | -99.139
4.0057 | -99.0929
4.0094 | -99.0069
4.0131 | -98.9041
4.0168 | -98.7955
4.0205 | -98.6827
4.0242 | -98.5745
4.0278 | -98.4793
4.0315 | -98.3954
4.0352 | -98.315
4.0389 | -98.2466
4.0426 | -98.1958
4.0463 | -98.1673
4.05 | -98.1733
4.0537 | -98.2155
4.0574 | -98.2787
4.061 | -98.3556
4.0647 | -98.4241
4.0684 | -98.4761
4.0721 | -98.5123
4.0758 | -98.5274
4.0795 | -98.5318
4.0832 | -98.5524
4.0869 | -98.6068
4.0906 | -98.6857
4.0942 | -98.7902
4.0979 | -98.9257
4.1016 | -99.0709
4.1053 | -99.2143
4.109 | -99.3485
4.1127 | -99.4521
4.1164 | -99.5314
4.1201 | -99.5838
4.1237 | -99.6093
4.1274 | -99.6159
4.1311 | -99.5894
4.1348 | -99.5395
4.1385 | -99.4812
4.1422 | -99.4297
4.1459 | -99.3953
4.1496 | -99.3844
4.1533 | -99.3999
4.1569 | -99.4357
4.1606 | -99.4867
4.1643 | -99.5462
4.168 | -99.6084
4.1717 | -99.6561
4.1754 | -99.6873
4.1791 | -99.6874
4.1828 | -99.6516
4.1865 | -99.5841
4.1901 | -99.51
4.1938 | -99.4405
4.1975 | -99.3815
4.2012 | -99.3313
4.2049 | -99.2904
4.2086 | -99.2559
4.2123 | -99.22
4.216 | -99.1863
4.2197 | -99.145
4.2233 | -99.0843
4.227 | -99.0081
4.2307 | -98.9243
4.2344 | -98.8458
4.2381 | -98.777
4.2418 | -98.7301
4.2455 | -98.7083
4.2492 | -98.7137
4.2528 | -98.7393
4.2565 | -98.7955
4.2602 | -98.841
4.2639 | -98.8956
4.2676 | -98.9598
4.2713 | -99.0259
4.275 | -99.0923
4.2787 | -99.1553
4.2824 | -99.2043
4.286 | -99.2412
4.2897 | -99.2631
4.2934 | -99.2708
4.2971 | -99.2698
4.3008 | -99.2561
4.3045 | -99.2277
4.3082 | -99.1892
4.3119 | -99.1434
4.3156 | -99.0947
4.3192 | -99.0446
4.3229 | -98.9972
4.3266 | -98.947
4.3303 | -98.8962
4.334 | -98.8519
4.3377 | -98.8176
4.3414 | -98.7908
4.3451 | -98.7731
4.3487 | -98.7721
4.3524 | -98.7976
4.3561 | -98.8382
4.3598 | -98.8889
4.3635 | -98.947
4.3672 | -99.0094
4.3709 | -99.0782
4.3746 | -99.1368
4.3783 | -99.1776
4.3819 | -99.2075
4.3856 | -99.2348
4.3893 | -99.2694
4.393 | -99.3102
4.3967 | -99.3631
4.4004 | -99.4255
4.4041 | -99.4896
4.4078 | -99.5559
4.4115 | -99.6286
4.4151 | -99.7109
4.4188 | -99.7959
4.4225 | -99.8773
4.4262 | -99.9613
4.4299 | -100.0419
4.4336 | -100.1043
4.4373 | -100.1385
4.441 | -100.1378
4.4446 | -100.1072
4.4483 | -100.0432
4.452 | -99.9361
4.4557 | -99.789
4.4594 | -99.6263
4.4631 | -99.4691
4.4668 | -99.3243
4.4705 | -99.2014
4.4742 | -99.113
4.4778 | -99.0633
4.4815 | -99.0475
4.4852 | -99.0621
4.4889 | -99.11
4.4926 | -99.1735
4.4963 | -99.2447
4.5 | -99.3265
4.5037 | -99.4042
4.5074 | -99.4743
4.511 | -99.5284
4.5147 | -99.5615
4.5184 | -99.5693
4.5221 | -99.558
4.5258 | -99.5359
4.5295 | -99.5046
4.5332 | -99.4743
4.5369 | -99.4608
4.5406 | -99.4633
4.5442 | -99.4713
4.5479 | -99.4801
4.5516 | -99.4912
4.5553 | -99.5003
4.559 | -99.4981
4.5627 | -99.4961
4.5664 | -99.4932
4.5701 | -99.4834
4.5737 | -99.4861
4.5774 | -99.508
4.5811 | -99.5545
4.5848 | -99.6145
4.5885 | -99.6758
4.5922 | -99.7462
4.5959 | -99.8091
4.5996 | -99.8577
4.6033 | -99.8859
4.6069 | -99.8857
4.6106 | -99.8691
4.6143 | -99.8453
4.618 | -99.8181
4.6217 | -99.7948
4.6254 | -99.7888
4.6291 | -99.8085
4.6328 | -99.8466
4.6365 | -99.9042
4.6401 | -99.9707
4.6438 | -100.037
4.6475 | -100.1007
4.6512 | -100.1583
4.6549 | -100.2069
4.6586 | -100.2428
4.6623 | -100.2537
4.666 | -100.2444
4.6696 | -100.2089
4.6733 | -100.1575
4.677 | -100.0962
4.6807 | -100.0313
4.6844 | -99.9738
4.6881 | -99.9353
4.6918 | -99.9202
4.6955 | -99.9268
4.6992 | -99.9567
4.7028 | -100.0001
4.7065 | -100.0521
4.7102 | -100.1128
4.7139 | -100.1609
4.7176 | -100.1906
4.7213 | -100.2061
4.725 | -100.2036
4.7287 | -100.185
4.7324 | -100.1458
4.736 | -100.1013
4.7397 | -100.0595
4.7434 | -100.0172
4.7471 | -99.9795
4.7508 | -99.9498
4.7545 | -99.9365
4.7582 | -99.9367
4.7619 | -99.9519
4.7655 | -99.9888
4.7692 | -100.0403
4.7729 | -100.0904
4.7766 | -100.1366
4.7803 | -100.1636
4.784 | -100.1747
4.7877 | -100.133
4.7914 | -100.1128
4.7951 | -100.0795
4.7987 | -100.0472
4.8024 | -100.0176
4.8061 | -99.9977
4.8098 | -99.9923
4.8135 | -100.0045
4.8172 | -100.0289
4.8209 | -100.0639
4.8246 | -100.0994
4.8283 | -100.1363
4.8319 | -100.1767
4.8356 | -100.2165
4.8393 | -100.2498
4.843 | -100.2784
4.8467 | -100.3104
4.8504 | -100.3348
4.8541 | -100.3593
4.8578 | -100.3801
4.8615 | -100.3976
4.8651 | -100.4126
4.8688 | -100.4282
4.8725 | -100.4457
4.8762 | -100.4517
4.8799 | -100.4502
4.8836 | -100.4426
4.8873 | -100.4275
4.891 | -100.4059
4.8946 | -100.3855
4.8983 | -100.3717
4.902 | -100.3688
4.9057 | -100.3674
4.9094 | -100.3715
4.9131 | -100.3857
4.9168 | -100.4053
4.9205 | -100.4334
4.9242 | -100.4633
4.9278 | -100.4949
4.9315 | -100.5169
4.9352 | -100.5289
4.9389 | -100.5345
4.9426 | -100.5285
4.9463 | -100.5084
4.95 | -100.4819
4.9537 | -100.4641
4.9574 | -100.4526
4.961 | -100.4423
4.9647 | -100.4418
4.9684 | -100.4506
4.9721 | -100.461
4.9758 | -100.4702
4.9795 | -100.484
4.9832 | -100.4989
4.9869 | -100.5215
4.9905 | -100.5449
4.9942 | -100.5677
4.9979 | -100.5897
5.0016 | -100.6267
5.0053 | -100.6688
5.009 | -100.7143
5.0127 | -100.7627
5.0164 | -100.7998
5.0201 | -100.8257
5.0237 | -100.8301
5.0274 | -100.8123
5.0311 | -100.7786
5.0348 | -100.7235
5.0385 | -100.6569
5.0422 | -100.5797
5.0459 | -100.5026
5.0496 | -100.4292
5.0533 | -100.3565
5.0569 | -100.2929
5.0606 | -100.2432
5.0643 | -100.1983
5.068 | -100.1713
5.0717 | -100.1638
5.0754 | -100.18
5.0791 | -100.2122
5.0828 | -100.2688
5.0864 | -100.3496
5.0901 | -100.4484
5.0938 | -100.5609
5.0975 | -100.6781
5.1012 | -100.7981
5.1049 | -100.9046
5.1086 | -100.9936
5.1123 | -101.0832
5.116 | -101.16
5.1196 | -101.2205
5.1233 | -101.2638
5.127 | -101.2976
5.1307 | -101.3204
5.1344 | -101.3295
5.1381 | -101.3331
5.1418 | -101.3433
5.1455 | -101.3412
5.1492 | -101.331
5.1528 | -101.3081
5.1565 | -101.2771
5.1602 | -101.2407
5.1639 | -101.2009
5.1676 | -101.1461
5.1713 | -101.0913
5.175 | -101.0507
5.1787 | -101.0123
5.1824 | -100.986
5.186 | -100.9728
5.1897 | -100.9636
5.1934 | -100.9677
5.1971 | -100.9727
5.2008 | -100.977
5.2045 | -100.9758
5.2082 | -100.9668
5.2119 | -100.9463
5.2155 | -100.9195
5.2192 | -100.8909
5.2229 | -100.8715
5.2266 | -100.858
5.2303 | -100.8558
5.234 | -100.8588
5.2377 | -100.8706
5.2414 | -100.8892
5.2451 | -100.9176
5.2487 | -100.9467
5.2524 | -100.9679
5.2561 | -100.9973
5.2598 | -101.0305
5.2635 | -101.0638
5.2672 | -101.0936
5.2709 | -101.1207
5.2746 | -101.1444
5.2783 | -101.1627
5.2819 | -101.1718
5.2856 | -101.1731
5.2893 | -101.1613
5.293 | -101.1376
5.2967 | -101.1038
5.3004 | -101.0651
5.3041 | -101.0238
5.3078 | -100.9839
5.3114 | -100.957
5.3151 | -100.9305
5.3188 | -100.9235
5.3225 | -100.9114
5.3262 | -100.899
5.3299 | -100.8779
5.3336 | -100.8593
5.3373 | -100.8388
5.341 | -100.8107
5.3446 | -100.7878
5.3483 | -100.7709
5.352 | -100.7544
5.3557 | -100.7439
5.3594 | -100.7325
5.3631 | -100.7288
5.3668 | -100.7378
5.3705 | -100.7607
5.3742 | -100.7979
5.3778 | -100.8485
5.3815 | -100.9177
5.3852 | -101.0059
5.3889 | -101.1029
5.3926 | -101.2028
5.3963 | -101.3042
5.4 | -101.3976
5.4037 | -101.4755
5.4073 | -101.5411
5.411 | -101.5919
5.4147 | -101.6346
5.4184 | -101.6496
5.4221 | -101.6443
5.4258 | -101.6231
5.4295 | -101.5792
5.4332 | -101.5234
5.4369 | -101.4643
5.4405 | -101.3949
5.4442 | -101.3308
5.4479 | -101.2708
5.4516 | -101.2089
5.4553 | -101.1553
5.459 | -101.1133
5.4627 | -101.0839
5.4664 | -101.0678
5.4701 | -101.0562
5.4737 | -101.0591
5.4774 | -101.0694
5.4811 | -101.0817
5.4848 | -101.0907
5.4885 | -101.0958
5.4922 | -101.0978
5.4959 | -101.0954
5.4996 | -101.0911
5.5033 | -101.0891
5.5069 | -101.0793
5.5106 | -101.0729
5.5143 | -101.0789
5.518 | -101.0934
5.5217 | -101.1107
5.5254 | -101.15
5.5291 | -101.2054
5.5328 | -101.2773
5.5364 | -101.3684
5.5401 | -101.471
5.5438 | -101.5726
5.5475 | -101.6684
5.5512 | -101.7439
5.5549 | -101.8105
5.5586 | -101.858
5.5623 | -101.8791
5.566 | -101.8732
5.5696 | -101.8542
5.5733 | -101.8205
5.577 | -101.7748
5.5807 | -101.7269
5.5844 | -101.6768
5.5881 | -101.6287
5.5918 | -101.5952
5.5955 | -101.5643
5.5992 | -101.5342
5.6028 | -101.5195
5.6065 | -101.5176
5.6102 | -101.5203
5.6139 | -101.5283
5.6176 | -101.538
5.6213 | -101.5519
5.625 | -101.5638
5.6287 | -101.5722
5.6323 | -101.571
5.636 | -101.5652
5.6397 | -101.553
5.6434 | -101.5326
5.6471 | -101.521
5.6508 | -101.509
5.6545 | -101.4798
5.6582 | -101.4545
5.6619 | -101.4252
5.6655 | -101.3932
5.6692 | -101.3562
5.6729 | -101.3237
5.6766 | -101.2885
5.6803 | -101.2518
5.684 | -101.2198
5.6877 | -101.1899
5.6914 | -101.1751
5.6951 | -101.1673
5.6987 | -101.1715
5.7024 | -101.1856
5.7061 | -101.2065
5.7098 | -101.2322
5.7135 | -101.2639
5.7172 | -101.3107
5.7209 | -101.3693
5.7246 | -101.4262
5.7282 | -101.4868
5.7319 | -101.5415
5.7356 | -101.5894
5.7393 | -101.6264
5.743 | -101.6391
5.7467 | -101.6278
5.7504 | -101.594
5.7541 | -101.5453
5.7578 | -101.4795
5.7614 | -101.4184
5.7651 | -101.3683
5.7688 | -101.3217
5.7725 | -101.293
5.7762 | -101.2794
5.7799 | -101.2821
5.7836 | -101.2946
5.7873 | -101.3195
5.791 | -101.3507
5.7946 | -101.3875
5.7983 | -101.4239
5.802 | -101.45
5.8057 | -101.4699
5.8094 | -101.4894
5.8131 | -101.5081
5.8168 | -101.524
5.8205 | -101.5353
5.8242 | -101.5394
5.8278 | -101.5436
5.8315 | -101.5574
5.8352 | -101.5767
5.8389 | -101.59
5.8426 | -101.5964
5.8463 | -101.5928
5.85 | -101.5508
5.8537 | -101.5214
5.8573 | -101.4894
5.861 | -101.448
5.8647 | -101.4046
5.8684 | -101.3619
5.8721 | -101.322
5.8758 | -101.2862
5.8795 | -101.2588
5.8832 | -101.2348
5.8869 | -101.2263
5.8905 | -101.2256
5.8942 | -101.2281
5.8979 | -101.2396
5.9016 | -101.2548
5.9053 | -101.2815
5.909 | -101.3164
5.9127 | -101.3577
5.9164 | -101.3949
5.9201 | -101.4303
5.9237 | -101.4607
5.9274 | -101.4951
5.9311 | -101.5321
5.9348 | -101.5647
5.9385 | -101.5943
5.9422 | -101.6262
5.9459 | -101.6552
5.9496 | -101.6802
5.9532 | -101.7023
5.9569 | -101.7203
5.9606 | -101.7415
5.9643 | -101.7615
5.968 | -101.7729
5.9717 | -101.7811
5.9754 | -101.7893
5.9791 | -101.7899
5.9828 | -101.7955
5.9864 | -101.8033
5.9901 | -101.8112
5.9938 | -101.8155
5.9975 | -101.8135
6.0012 | -101.8137
6.0049 | -101.8115
6.0086 | -101.8089
6.0123 | -101.8007
6.016 | -101.7917
6.0196 | -101.7749
6.0233 | -101.7566
6.027 | -101.7438
6.0307 | -101.7348
6.0344 | -101.7245
6.0381 | -101.7236
6.0418 | -101.7323
6.0455 | -101.7468
6.0491 | -101.766
6.0528 | -101.7847
6.0565 | -101.8038
6.0602 | -101.8295
6.0639 | -101.8574
6.0676 | -101.8882
6.0713 | -101.9172
6.075 | -101.9388
6.0787 | -101.9579
6.0823 | -101.9644
6.086 | -101.9636
6.0897 | -101.9583
6.0934 | -101.9405
6.0971 | -101.9168
6.1008 | -101.891
6.1045 | -101.8648
6.1082 | -101.835
6.1119 | -101.8071
6.1155 | -101.7763
6.1192 | -101.7396
6.1229 | -101.7031
6.1266 | -101.6672
6.1303 | -101.6311
6.134 | -101.5974
6.1377 | -101.5734
6.1414 | -101.558
6.1451 | -101.5508
6.1487 | -101.5516
6.1524 | -101.5683
6.1561 | -101.5996
6.1598 | -101.6404
6.1635 | -101.6877
6.1672 | -101.7428
6.1709 | -101.7991
6.1746 | -101.8558
6.1782 | -101.9085
6.1819 | -101.9475
6.1856 | -101.973
6.1893 | -101.9836
6.193 | -101.981
6.1967 | -101.9741
6.2004 | -101.9615
6.2041 | -101.9443
6.2078 | -101.9265
6.2114 | -101.9207
6.2151 | -101.9259
6.2188 | -101.9405
6.2225 | -101.9603
6.2262 | -101.9943
6.2299 | -102.036
6.2336 | -102.0892
6.2373 | -102.1489
6.241 | -102.2183
6.2446 | -102.2791
6.2483 | -102.3295
6.252 | -102.3763
6.2557 | -102.4207
6.2594 | -102.4526
6.2631 | -102.4691
6.2668 | -102.4764
6.2705 | -102.4712
6.2741 | -102.4586
6.2778 | -102.4452
6.2815 | -102.4265
6.2852 | -102.3953
6.2889 | -102.3552
6.2926 | -102.3246
6.2963 | -102.2943
6.3 | -102.2614
6.3037 | -102.2315
6.3073 | -102.2123
6.311 | -102.1996
6.3147 | -102.1947
6.3184 | -102.1817
6.3221 | -102.1696
6.3258 | -102.158
6.3295 | -102.1445
6.3332 | -102.1219
6.3369 | -102.0952
6.3405 | -102.065
6.3442 | -102.0317
6.3479 | -101.9947
6.3516 | -101.9602
6.3553 | -101.9338
6.359 | -101.906
6.3627 | -101.887
6.3664 | -101.8782
6.37 | -101.8815
6.3737 | -101.8998
6.3774 | -101.9349
6.3811 | -101.9818
6.3848 | -102.051
6.3885 | -102.1046
6.3922 | -102.1653
6.3959 | -102.2261
6.3996 | -102.2784
6.4032 | -102.3201
6.4069 | -102.3603
6.4106 | -102.3956
6.4143 | -102.4246
6.418 | -102.4407
6.4217 | -102.4459
6.4254 | -102.4421
6.4291 | -102.4391
6.4328 | -102.4373
6.4364 | -102.4357
6.4401 | -102.4311
6.4438 | -102.4292
6.4475 | -102.4263
6.4512 | -102.4236
6.4549 | -102.4286
6.4586 | -102.4332
6.4623 | -102.4358
6.466 | -102.4369
6.4696 | -102.4375
6.4733 | -102.4422
6.477 | -102.4454
6.4807 | -102.4467
6.4844 | -102.4517
6.4881 | -102.4576
6.4918 | -102.4588
6.4955 | -102.462
6.4991 | -102.4681
6.5028 | -102.4662
6.5065 | -102.4512
6.5102 | -102.439
6.5139 | -102.4299
6.5176 | -102.4192
6.5213 | -102.401
6.525 | -102.3848
6.5287 | -102.3663
6.5323 | -102.3464
6.536 | -102.325
6.5397 | -102.2982
6.5434 | -102.2664
6.5471 | -102.2389
6.5508 | -102.2075
6.5545 | -102.1719
6.5582 | -102.1375
6.5619 | -102.1045
6.5655 | -102.0675
6.5692 | -102.0303
6.5729 | -101.9981
6.5766 | -101.9742
6.5803 | -101.9594
6.584 | -101.9498
6.5877 | -101.9554
6.5914 | -101.9746
6.595 | -102.0037
6.5987 | -102.0425
6.6024 | -102.1008
6.6061 | -102.171
6.6098 | -102.2367
6.6135 | -102.3017
6.6172 | -102.3643
6.6209 | -102.4294
6.6246 | -102.4841
6.6282 | -102.5246
6.6319 | -102.5523
6.6356 | -102.5696
6.6393 | -102.5774
6.643 | -102.5723
6.6467 | -102.5569
6.6504 | -102.5361
6.6541 | -102.5133
6.6578 | -102.492
6.6614 | -102.47
6.6651 | -102.4543
6.6688 | -102.4412
6.6725 | -102.4216
6.6762 | -102.4024
6.6799 | -102.384
6.6836 | -102.372
6.6873 | -102.3674
6.6909 | -102.3719
6.6946 | -102.3801
6.6983 | -102.3942
6.702 | -102.414
6.7057 | -102.4405
6.7094 | -102.4729
6.7131 | -102.5087
6.7168 | -102.5374
6.7205 | -102.5539
6.7241 | -102.5713
6.7278 | -102.5817
6.7315 | -102.5862
6.7352 | -102.5799
6.7389 | -102.5622
6.7426 | -102.5372
6.7463 | -102.5076
6.75 | -102.4752
6.7537 | -102.4414
6.7573 | -102.4122
6.761 | -102.3936
6.7647 | -102.3808
6.7684 | -102.3812
6.7721 | -102.3953
6.7758 | -102.4198
6.7795 | -102.4536
6.7832 | -102.4907
6.7869 | -102.5253
6.7905 | -102.5597
6.7942 | -102.593
6.7979 | -102.6152
6.8016 | -102.6273
6.8053 | -102.6298
6.809 | -102.6242
6.8127 | -102.6049
6.8164 | -102.5809
6.82 | -102.5565
6.8237 | -102.5264
6.8274 | -102.499
6.8311 | -102.4764
6.8348 | -102.4638
6.8385 | -102.4593
6.8422 | -102.463
6.8459 | -102.472
6.8496 | -102.4901
6.8532 | -102.5104
6.8569 | -102.5316
6.8606 | -102.5611
6.8643 | -102.6002
6.868 | -102.6406
6.8717 | -102.6858
6.8754 | -102.7296
6.8791 | -102.7743
6.8828 | -102.8201
6.8864 | -102.8658
6.8901 | -102.9104
6.8938 | -102.9537
6.8975 | -102.9932
6.9012 | -103.0295
6.9049 | -103.0602
6.9086 | -103.0896
6.9123 | -103.1235
6.9159 | -103.1474
6.9196 | -103.1694
6.9233 | -103.1816
6.927 | -103.1851
6.9307 | -103.1831
6.9344 | -103.1709
6.9381 | -103.1538
6.9418 | -103.1241
6.9455 | -103.0898
6.9491 | -103.0502
6.9528 | -103.0082
6.9565 | -102.9736
6.9602 | -102.9404
6.9639 | -102.9081
6.9676 | -102.8769
6.9713 | -102.8506
6.975 | -102.8353
6.9787 | -102.8283
6.9823 | -102.8207
6.986 | -102.819
6.9897 | -102.8167
6.9934 | -102.8215
6.9971 | -102.8308
7.0008 | -102.8416
7.0045 | -102.8497
7.0082 | -102.8583
7.0118 | -102.8685
7.0155 | -102.875
7.0192 | -102.8828
7.0229 | -102.8947
7.0266 | -102.9151
7.0303 | -102.9389
7.034 | -102.9628
7.0377 | -102.9868
7.0414 | -103.0145
7.045 | -103.0337
7.0487 | -103.0537
7.0524 | -103.0689
7.0561 | -103.0814
7.0598 | -103.0882
7.0635 | -103.0843
7.0672 | -103.0839
7.0709 | -103.0839
7.0746 | -103.0891
7.0782 | -103.0949
7.0819 | -103.096
7.0856 | -103.0996
7.0893 | -103.1078
7.093 | -103.1196
7.0967 | -103.1377
7.1004 | -103.1471
7.1041 | -103.1552
7.1078 | -103.1668
7.1114 | -103.1811
7.1151 | -103.1903
7.1188 | -103.1885
7.1225 | -103.1867
7.1262 | -103.1849
7.1299 | -103.1759
7.1336 | -103.1602
7.1373 | -103.136
7.1409 | -103.1103
7.1446 | -103.0788
7.1483 | -103.0381
7.152 | -102.9979
7.1557 | -102.9562
7.1594 | -102.9097
7.1631 | -102.8587
7.1668 | -102.8029
7.1705 | -102.7601
7.1741 | -102.7177
7.1778 | -102.6809
7.1815 | -102.6528
7.1852 | -102.6382
7.1889 | -102.6347
7.1926 | -102.6378
7.1963 | -102.6488
7.2 | -102.6632
7.2037 | -102.6873
7.2073 | -102.723
7.211 | -102.7652
7.2147 | -102.8108
7.2184 | -102.861
7.2221 | -102.9178
7.2258 | -102.9815
7.2295 | -103.0437
7.2332 | -103.1031
7.2368 | -103.1629
7.2405 | -103.2227
7.2442 | -103.2791
7.2479 | -103.3216
7.2516 | -103.3556
7.2553 | -103.3807
7.259 | -103.3971
7.2627 | -103.4069
7.2664 | -103.4107
7.27 | -103.4133
7.2737 | -103.4145
7.2774 | -103.3997
7.2811 | -103.3876
7.2848 | -103.372
7.2885 | -103.3511
7.2922 | -103.3285
7.2959 | -103.311
7.2996 | -103.2976
7.3032 | -103.2797
7.3069 | -103.2607
7.3106 | -103.2427
7.3143 | -103.2241
7.318 | -103.2076
7.3217 | -103.1953
7.3254 | -103.1879
7.3291 | -103.1822
7.3327 | -103.1728
7.3364 | -103.1659
7.3401 | -103.1652
7.3438 | -103.1655
7.3475 | -103.168
7.3512 | -103.1747
7.3549 | -103.1806
7.3586 | -103.1824
7.3623 | -103.1784
7.3659 | -103.1647
7.3696 | -103.1421
7.3733 | -103.1096
7.377 | -103.0698
7.3807 | -103.0285
7.3844 | -102.9841
7.3881 | -102.9373
7.3918 | -102.9
7.3955 | -102.8668
7.3991 | -102.8295
7.4028 | -102.7879
7.4065 | -102.757
7.4102 | -102.7398
7.4139 | -102.7287
7.4176 | -102.7224
7.4213 | -102.7223
7.425 | -102.7318
7.4287 | -102.7441
7.4323 | -102.7649
7.436 | -102.7889
7.4397 | -102.816
7.4434 | -102.8462
7.4471 | -102.8724
7.4508 | -102.8987
7.4545 | -102.9287
7.4582 | -102.9492
7.4618 | -102.964
7.4655 | -102.9825
7.4692 | -102.9958
7.4729 | -103.0071
7.4766 | -103.0175
7.4803 | -103.0242
7.484 | -103.0275
7.4877 | -103.03
7.4914 | -103.0302
7.495 | -103.0378
7.4987 | -103.0459
7.5024 | -103.0518
7.5061 | -103.0659
7.5098 | -103.0915
7.5135 | -103.121
7.5172 | -103.1455
7.5209 | -103.1736
7.5246 | -103.2031
7.5282 | -103.2322
7.5319 | -103.2654
7.5356 | -103.295
7.5393 | -103.3274
7.543 | -103.3491
7.5467 | -103.3732
7.5504 | -103.3966
7.5541 | -103.4119
7.5577 | -103.4239
7.5614 | -103.4319
7.5651 | -103.4319
7.5688 | -103.4287
7.5725 | -103.4266
7.5762 | -103.4186
7.5799 | -103.4108
7.5836 | -103.3989
7.5873 | -103.3781
7.5909 | -103.354
7.5946 | -103.3397
7.5983 | -103.3273
7.602 | -103.3159
7.6057 | -103.3025
7.6094 | -103.2814
7.6131 | -103.2638
7.6168 | -103.2493
7.6205 | -103.2324
7.6241 | -103.217
7.6278 | -103.2073
7.6315 | -103.1977
7.6352 | -103.1837
7.6389 | -103.1789
7.6426 | -103.1796
7.6463 | -103.1868
7.65 | -103.199
7.6536 | -103.2152
7.6573 | -103.2433
7.661 | -103.2727
7.6647 | -103.3069
7.6684 | -103.3415
7.6721 | -103.3779
7.6758 | -103.4203
7.6795 | -103.4633
7.6832 | -103.497
7.6868 | -103.5315
7.6905 | -103.5646
7.6942 | -103.5881
7.6979 | -103.6114
7.7016 | -103.6328
7.7053 | -103.6446
7.709 | -103.6425
7.7127 | -103.647
7.7164 | -103.6423
7.72 | -103.6344
7.7237 | -103.6303
7.7274 | -103.6244
7.7311 | -103.6171
7.7348 | -103.6075
7.7385 | -103.5929
7.7422 | -103.5769
7.7459 | -103.5632
7.7496 | -103.5545
7.7532 | -103.5427
7.7569 | -103.5339
7.7606 | -103.5252
7.7643 | -103.5187
7.768 | -103.5165
7.7717 | -103.5131
7.7754 | -103.5172
7.7791 | -103.5252
7.7827 | -103.5346
7.7864 | -103.5373
7.7901 | -103.5347
7.7938 | -103.5323
7.7975 | -103.5343
7.8012 | -103.5297
7.8049 | -103.5225
7.8086 | -103.5135
7.8123 | -103.5076
7.8159 | -103.501
7.8196 | -103.493
7.8233 | -103.4832
7.827 | -103.4705
7.8307 | -103.4666
7.8344 | -103.4683
7.8381 | -103.4691
7.8418 | -103.4684
7.8455 | -103.4705
7.8491 | -103.4765
7.8528 | -103.4853
7.8565 | -103.4926
7.8602 | -103.4996
7.8639 | -103.5035
7.8676 | -103.5078
7.8713 | -103.5152
7.875 | -103.5156
7.8786 | -103.5075
7.8823 | -103.4948
7.886 | -103.478
7.8897 | -103.4556
7.8934 | -103.4296
7.8971 | -103.4051
7.9008 | -103.3856
7.9045 | -103.3653
7.9082 | -103.3503
7.9118 | -103.3384
7.9155 | -103.3338
7.9192 | -103.336
7.9229 | -103.348
7.9266 | -103.3642
7.9303 | -103.3853
7.934 | -103.412
7.9377 | -103.4405
7.9414 | -103.4639
7.945 | -103.4865
7.9487 | -103.5079
7.9524 | -103.5269
7.9561 | -103.5493
7.9598 | -103.5684
7.9635 | -103.5799
7.9672 | -103.5905
7.9709 | -103.6009
7.9745 | -103.607
7.9782 | -103.6075
7.9819 | -103.6004
7.9856 | -103.5972
7.9893 | -103.588
7.993 | -103.5806
7.9967 | -103.5807
8.0004 | -103.577
8.0041 | -103.5747
8.0077 | -103.5746
8.0114 | -103.5778
8.0151 | -103.5893
8.0188 | -103.6027
8.0225 | -103.6165
8.0262 | -103.6366
8.0299 | -103.6632
8.0336 | -103.688
8.0373 | -103.7154
8.0409 | -103.7414
8.0446 | -103.7651
8.0483 | -103.7884
8.052 | -103.8076
8.0557 | -103.8204
8.0594 | -103.8301
8.0631 | -103.8349
8.0668 | -103.8326
8.0704 | -103.8303
8.0741 | -103.8229
8.0778 | -103.8129
8.0815 | -103.8048
8.0852 | -103.7921
8.0889 | -103.7746
8.0926 | -103.7559
8.0963 | -103.7406
8.1 | -103.7279
8.1036 | -103.7219
8.1073 | -103.7199
8.111 | -103.7193
8.1147 | -103.7216
8.1184 | -103.7207
8.1221 | -103.7291
8.1258 | -103.7371
8.1295 | -103.7474
8.1332 | -103.7598
8.1368 | -103.7684
8.1405 | -103.7807
8.1442 | -103.7986
8.1479 | -103.8083
8.1516 | -103.8165
8.1553 | -103.8266
8.159 | -103.8355
8.1627 | -103.8433
8.1664 | -103.8502
8.17 | -103.8549
8.1737 | -103.8588
8.1774 | -103.8596
8.1811 | -103.8591
8.1848 | -103.8574
8.1885 | -103.8525
8.1922 | -103.8428
8.1959 | -103.8319
8.1995 | -103.8263
8.2032 | -103.8223
8.2069 | -103.8218
8.2106 | -103.8246
8.2143 | -103.8288
8.218 | -103.8345
8.2217 | -103.84
8.2254 | -103.8391
8.2291 | -103.8386
8.2327 | -103.8411
8.2364 | -103.8406
8.2401 | -103.8389
8.2438 | -103.8314
8.2475 | -103.825
8.2512 | -103.8137
8.2549 | -103.7959
8.2586 | -103.7862
8.2623 | -103.7745
8.2659 | -103.763
8.2696 | -103.7533
8.2733 | -103.7475
8.277 | -103.7459
8.2807 | -103.7472
8.2844 | -103.7549
8.2881 | -103.7658
8.2918 | -103.7758
8.2954 | -103.7898
8.2991 | -103.8067
8.3028 | -103.8234
8.3065 | -103.8359
8.3102 | -103.8491
8.3139 | -103.8608
8.3176 | -103.8661
8.3213 | -103.8601
8.325 | -103.8493
8.3286 | -103.8298
8.3323 | -103.8056
8.336 | -103.7766
8.3397 | -103.7458
8.3434 | -103.7174
8.3471 | -103.6947
8.3508 | -103.6759
8.3545 | -103.6624
8.3582 | -103.6582
8.3618 | -103.6562
8.3655 | -103.6592
8.3692 | -103.6698
8.3729 | -103.6908
8.3766 | -103.7184
8.3803 | -103.7395
8.384 | -103.766
8.3877 | -103.7916
8.3913 | -103.8198
8.395 | -103.8455
8.3987 | -103.8691
8.4024 | -103.8889
8.4061 | -103.904
8.4098 | -103.9145
8.4135 | -103.9229
8.4172 | -103.9215
8.4209 | -103.9199
8.4245 | -103.9148
8.4282 | -103.9041
8.4319 | -103.8875
8.4356 | -103.8692
8.4393 | -103.8471
8.443 | -103.8204
8.4467 | -103.7931
8.4504 | -103.7657
8.4541 | -103.7389
8.4577 | -103.7113
8.4614 | -103.6824
8.4651 | -103.6522
8.4688 | -103.6241
8.4725 | -103.5964
8.4762 | -103.5756
8.4799 | -103.5602
8.4836 | -103.5534
8.4873 | -103.5512
8.4909 | -103.5552
8.4946 | -103.5675
8.4983 | -103.5879
8.502 | -103.6186
8.5057 | -103.6564
8.5094 | -103.7007
8.5131 | -103.754
8.5168 | -103.7993
8.5204 | -103.8454
8.5241 | -103.8926
8.5278 | -103.9394
8.5315 | -103.984
8.5352 | -104.0266
8.5389 | -104.0641
8.5426 | -104.0941
8.5463 | -104.1121
8.55 | -104.125
8.5536 | -104.1342
8.5573 | -104.1353
8.561 | -104.1318
8.5647 | -104.1246
8.5684 | -104.1113
8.5721 | -104.0888
8.5758 | -104.0707
8.5795 | -104.0443
8.5832 | -104.0126
8.5868 | -103.9823
8.5905 | -103.9552
8.5942 | -103.9295
8.5979 | -103.9044
8.6016 | -103.8699
8.6053 | -103.8367
8.609 | -103.8055
8.6127 | -103.7801
8.6163 | -103.7596
8.62 | -103.7405
8.6237 | -103.7314
8.6274 | -103.729
8.6311 | -103.7295
8.6348 | -103.741
8.6385 | -103.763
8.6422 | -103.7877
8.6459 | -103.8212
8.6495 | -103.8665
8.6532 | -103.9114
8.6569 | -103.9597
8.6606 | -104.012
8.6643 | -104.0619
8.668 | -104.1105
8.6717 | -104.1598
8.6754 | -104.2044
8.6791 | -104.241
8.6827 | -104.2738
8.6864 | -104.2972
8.6901 | -104.3051
8.6938 | -104.3043
8.6975 | -104.2952
8.7012 | -104.272
8.7049 | -104.236
8.7086 | -104.2035
8.7122 | -104.1683
8.7159 | -104.1217
8.7196 | -104.0721
8.7233 | -104.0275
8.727 | -103.9811
8.7307 | -103.944
8.7344 | -103.9063
8.7381 | -103.8678
8.7418 | -103.8341
8.7454 | -103.8108
8.7491 | -103.7921
8.7528 | -103.7741
8.7565 | -103.7651
8.7602 | -103.7619
8.7639 | -103.7565
8.7676 | -103.7494
8.7713 | -103.7469
8.775 | -103.741
8.7786 | -103.7341
8.7823 | -103.7323
8.786 | -103.7285
8.7897 | -103.7245
8.7934 | -103.7211
8.7971 | -103.7145
8.8008 | -103.7082
8.8045 | -103.7048
8.8082 | -103.701
8.8118 | -103.6959
8.8155 | -103.6946
8.8192 | -103.6964
8.8229 | -103.6991
8.8266 | -103.7036
8.8303 | -103.7148
8.834 | -103.7247
8.8377 | -103.7351
8.8413 | -103.7544
8.845 | -103.7803
8.8487 | -103.8046
8.8524 | -103.8293
8.8561 | -103.861
8.8598 | -103.8884
8.8635 | -103.9145
8.8672 | -103.9358
8.8709 | -103.9517
8.8745 | -103.9616
8.8782 | -103.969
8.8819 | -103.9759
8.8856 | -103.9735
8.8893 | -103.9671
8.893 | -103.9583
8.8967 | -103.9525
8.9004 | -103.9455
8.9041 | -103.936
8.9077 | -103.925
8.9114 | -103.9163
8.9151 | -103.914
8.9188 | -103.9166
8.9225 | -103.9235
8.9262 | -103.9344
8.9299 | -103.94
8.9336 | -103.9513
8.9372 | -103.9639
8.9409 | -103.979
8.9446 | -103.9977
8.9483 | -104.01
8.952 | -104.0219
8.9557 | -104.0331
8.9594 | -104.0376
8.9631 | -104.0403
8.9668 | -104.0419
8.9704 | -104.0419
8.9741 | -104.0463
8.9778 | -104.0575
8.9815 | -104.0703
8.9852 | -104.0839
8.9889 | -104.1032
8.9926 | -104.1278
8.9963 | -104.1559
9 | -104.1807
9.0036 | -104.2049
9.0073 | -104.2319
9.011 | -104.2586
9.0147 | -104.2911
9.0184 | -104.3197
9.0221 | -104.3466
9.0258 | -104.3706
9.0295 | -104.3889
9.0331 | -104.4038
9.0368 | -104.4197
9.0405 | -104.4324
9.0442 | -104.4305
9.0479 | -104.4308
9.0516 | -104.4279
9.0553 | -104.4315
9.059 | -104.4328
9.0627 | -104.4335
9.0663 | -104.4256
9.07 | -104.4136
9.0737 | -104.4008
9.0774 | -104.3892
9.0811 | -104.3808
9.0848 | -104.3643
9.0885 | -104.3458
9.0922 | -104.3304
9.0959 | -104.3151
9.0995 | -104.2981
9.1032 | -104.2796
9.1069 | -104.2577
9.1106 | -104.2382
9.1143 | -104.2214
9.118 | -104.2028
9.1217 | -104.1866
9.1254 | -104.1724
9.1291 | -104.1632
9.1327 | -104.161
9.1364 | -104.162
9.1401 | -104.1655
9.1438 | -104.1704
9.1475 | -104.1739
9.1512 | -104.1893
9.1549 | -104.2069
9.1586 | -104.2238
9.1622 | -104.2438
9.1659 | -104.2706
9.1696 | -104.2953
9.1733 | -104.3197
9.177 | -104.3431
9.1807 | -104.366
9.1844 | -104.3926
9.1881 | -104.4165
9.1918 | -104.4339
9.1954 | -104.4529
9.1991 | -104.4764
9.2028 | -104.4978
9.2065 | -104.5199
9.2102 | -104.5361
9.2139 | -104.5475
9.2176 | -104.5586
9.2213 | -104.5758
9.225 | -104.5865
9.2286 | -104.5948
9.2323 | -104.6037
9.236 | -104.6109
9.2397 | -104.6149
9.2434 | -104.6229
9.2471 | -104.6249
9.2508 | -104.6286
9.2545 | -104.6347
9.2581 | -104.6342
9.2618 | -104.6316
9.2655 | -104.6253
9.2692 | -104.6188
9.2729 | -104.6125
9.2766 | -104.6103
9.2803 | -104.6063
9.284 | -104.6036
9.2877 | -104.5946
9.2913 | -104.585
9.295 | -104.5768
9.2987 | -104.5708
9.3024 | -104.5645
9.3061 | -104.5607
9.3098 | -104.5543
9.3135 | -104.5525
9.3172 | -104.5528
9.3209 | -104.548
9.3245 | -104.542
9.3282 | -104.5358
9.3319 | -104.5268
9.3356 | -104.5179
9.3393 | -104.5124
9.343 | -104.5092
9.3467 | -104.5027
9.3504 | -104.4916
9.3541 | -104.481
9.3577 | -104.4718
9.3614 | -104.4619
9.3651 | -104.4464
9.3688 | -104.4328
9.3725 | -104.4226
9.3762 | -104.4129
9.3799 | -104.4004
9.3836 | -104.3921
9.3872 | -104.3828
9.3909 | -104.3795
9.3946 | -104.3762
9.3983 | -104.3723
9.402 | -104.3698
9.4057 | -104.3632
9.4094 | -104.3574
9.4131 | -104.353
9.4168 | -104.3541
9.4204 | -104.3524
9.4241 | -104.3544
9.4278 | -104.3562
9.4315 | -104.3581
9.4352 | -104.3602
9.4389 | -104.3622
9.4426 | -104.3621
9.4463 | -104.3692
9.45 | -104.3746
9.4536 | -104.3796
9.4573 | -104.3908
9.461 | -104.4008
9.4647 | -104.4094
9.4684 | -104.42
9.4721 | -104.4265
9.4758 | -104.431
9.4795 | -104.4294
9.4831 | -104.4268
9.4868 | -104.4257
9.4905 | -104.4195
9.4942 | -104.4065
9.4979 | -104.3956
9.5016 | -104.3868
9.5053 | -104.3759
9.509 | -104.366
9.5127 | -104.3576
9.5163 | -104.3547
9.52 | -104.3522
9.5237 | -104.3495
9.5274 | -104.3495
9.5311 | -104.3569
9.5348 | -104.3657
9.5385 | -104.3734
9.5422 | -104.382
9.5459 | -104.3917
9.5495 | -104.3979
9.5532 | -104.4033
9.5569 | -104.4126
9.5606 | -104.4189
9.5643 | -104.4221
9.568 | -104.4242
9.5717 | -104.4236
9.5754 | -104.413
9.579 | -104.413
9.5827 | -104.4122
9.5864 | -104.4097
9.5901 | -104.4061
9.5938 | -104.3989
9.5975 | -104.3877
9.6012 | -104.3785
9.6049 | -104.3697
9.6086 | -104.3557
9.6122 | -104.3436
9.6159 | -104.3332
9.6196 | -104.3283
9.6233 | -104.3271
9.627 | -104.3283
9.6307 | -104.3311
9.6344 | -104.3367
9.6381 | -104.3438
9.6418 | -104.3554
9.6454 | -104.3694
9.6491 | -104.3899
9.6528 | -104.4167
9.6565 | -104.4405
9.6602 | -104.4668
9.6639 | -104.4934
9.6676 | -104.5235
9.6713 | -104.5524
9.675 | -104.5819
9.6786 | -104.6132
9.6823 | -104.6454
9.686 | -104.6758
9.6897 | -104.7021
9.6934 | -104.7255
9.6971 | -104.7479
9.7008 | -104.7683
9.7045 | -104.791
9.7081 | -104.8143
9.7118 | -104.8352
9.7155 | -104.8557
9.7192 | -104.874
9.7229 | -104.8859
9.7266 | -104.9001
9.7303 | -104.9094
9.734 | -104.9184
9.7377 | -104.9285
9.7413 | -104.9381
9.745 | -104.9448
9.7487 | -104.9501
9.7524 | -104.9575
9.7561 | -104.959
9.7598 | -104.959
9.7635 | -104.9579
9.7672 | -104.9585
9.7709 | -104.9543
9.7745 | -104.9541
9.7782 | -104.957
9.7819 | -104.9634
9.7856 | -104.9707
9.7893 | -104.9807
9.793 | -104.9891
9.7967 | -105.0006
9.8004 | -105.0154
9.804 | -105.0287
9.8077 | -105.0444
9.8114 | -105.0535
9.8151 | -105.0625
9.8188 | -105.0645
9.8225 | -105.0567
9.8262 | -105.0439
9.8299 | -105.0292
9.8336 | -105.0105
9.8372 | -104.9874
9.8409 | -104.9562
9.8446 | -104.9215
9.8483 | -104.8825
9.852 | -104.8343
9.8557 | -104.7889
9.8594 | -104.74
9.8631 | -104.6881
9.8668 | -104.6382
9.8704 | -104.591
9.8741 | -104.5472
9.8778 | -104.502
9.8815 | -104.4605
9.8852 | -104.4261
9.8889 | -104.3985
9.8926 | -104.3749
9.8963 | -104.358
9.8999 | -104.3528
9.9036 | -104.3496
9.9073 | -104.3502
9.911 | -104.3573
9.9147 | -104.3711
9.9184 | -104.3867
9.9221 | -104.4017
9.9258 | -104.4202
9.9295 | -104.44
9.9331 | -104.4588
9.9368 | -104.4807
9.9405 | -104.4995
9.9442 | -104.5193
9.9479 | -104.537
9.9516 | -104.5485
9.9553 | -104.5564
9.959 | -104.5619
9.9627 | -104.5676
9.9663 | -104.5715
9.97 | -104.5694
9.9737 | -104.5602
9.9774 | -104.5476
9.9811 | -104.5318
9.9848 | -104.516
9.9885 | -104.5001
9.9922 | -104.4857
9.9959 | -104.4755
9.9995 | -104.4628
10.0032 | -104.4573
10.0069 | -104.4589
10.0106 | -104.4592
10.0143 | -104.4561
10.018 | -104.4611
10.0217 | -104.4713
10.0254 | -104.4844
10.029 | -104.4996
10.0327 | -104.5144
10.0364 | -104.5301
10.0401 | -104.5498
10.0438 | -104.5685
10.0475 | -104.593
10.0512 | -104.6226
10.0549 | -104.6466
10.0586 | -104.6708
10.0622 | -104.6953
10.0659 | -104.7209
10.0696 | -104.7435
10.0733 | -104.7607
10.077 | -104.7754
10.0807 | -104.7862
10.0844 | -104.7935
10.0881 | -104.7945
10.0918 | -104.8009
10.0954 | -104.8074
10.0991 | -104.8077
10.1028 | -104.8079
10.1065 | -104.8069
10.1102 | -104.8007
10.1139 | -104.7974
10.1176 | -104.7967
10.1213 | -104.7965
10.1249 | -104.7989
10.1286 | -104.7997
10.1323 | -104.8007
10.136 | -104.803
10.1397 | -104.8051
10.1434 | -104.8106
10.1471 | -104.8194
10.1508 | -104.8261
10.1545 | -104.8351
10.1581 | -104.8511
10.1618 | -104.8674
10.1655 | -104.8808
10.1692 | -104.898
10.1729 | -104.9134
10.1766 | -104.9244
10.1803 | -104.9373
10.184 | -104.9486
10.1877 | -104.9566
10.1913 | -104.9613
10.195 | -104.9601
10.1987 | -104.954
10.2024 | -104.9453
10.2061 | -104.9327
10.2098 | -104.9146
10.2135 | -104.89
10.2172 | -104.8635
10.2208 | -104.8371
10.2245 | -104.8108
10.2282 | -104.7824
10.2319 | -104.7603
10.2356 | -104.7395
10.2393 | -104.7147
10.243 | -104.6911
10.2467 | -104.6748
10.2504 | -104.6599
10.254 | -104.6479
10.2577 | -104.6416
10.2614 | -104.6365
10.2651 | -104.6363
10.2688 | -104.6356
10.2725 | -104.6369
10.2762 | -104.6398
10.2799 | -104.6441
10.2836 | -104.6503
10.2872 | -104.6581
10.2909 | -104.6628
10.2946 | -104.6652
10.2983 | -104.6716
10.302 | -104.6787
10.3057 | -104.6849
10.3094 | -104.6953
10.3131 | -104.7085
10.3168 | -104.7226
10.3204 | -104.7351
10.3241 | -104.7461
10.3278 | -104.7598
10.3315 | -104.7738
10.3352 | -104.7924
10.3389 | -104.8058
10.3426 | -104.817
10.3463 | -104.8295
10.3499 | -104.8419
10.3536 | -104.8571
10.3573 | -104.87
10.361 | -104.8813
10.3647 | -104.8931
10.3684 | -104.9052
10.3721 | -104.9167
10.3758 | -104.9276
10.3795 | -104.937
10.3831 | -104.9423
10.3868 | -104.9513
10.3905 | -104.9592
10.3942 | -104.9683
10.3979 | -104.978
10.4016 | -104.9847
10.4053 | -104.9915
10.409 | -104.9991
10.4127 | -105.009
10.4163 | -105.0139
10.42 | -105.0143
10.4237 | -105.0131
10.4274 | -105.0183
10.4311 | -105.0224
10.4348 | -105.0225
10.4385 | -105.0193
10.4422 | -105.018
10.4458 | -105.0168
10.4495 | -105.0074
10.4532 | -105.0036
10.4569 | -105.0007
10.4606 | -104.9914
10.4643 | -104.9812
10.468 | -104.9697
10.4717 | -104.9605
10.4754 | -104.9543
10.479 | -104.9455
10.4827 | -104.9297
10.4864 | -104.9161
10.4901 | -104.9055
10.4938 | -104.8966
10.4975 | -104.8904
10.5012 | -104.8825
10.5049 | -104.875
10.5086 | -104.8701
10.5122 | -104.8717
10.5159 | -104.8748
10.5196 | -104.8805
10.5233 | -104.8915
10.527 | -104.9025
10.5307 | -104.9211
10.5344 | -104.9421
10.5381 | -104.9594
10.5417 | -104.9803
10.5454 | -105.0019
10.5491 | -105.0201
10.5528 | -105.0391
10.5565 | -105.0591
10.5602 | -105.0712
10.5639 | -105.0777
10.5676 | -105.0832
10.5713 | -105.0895
10.5749 | -105.0898
10.5786 | -105.0796
10.5823 | -105.0695
10.586 | -105.0578
10.5897 | -105.0439
10.5934 | -105.0254
10.5971 | -105.0068
10.6008 | -104.9895
10.6045 | -104.9591
10.6081 | -104.9319
10.6118 | -104.9099
10.6155 | -104.8897
10.6192 | -104.8717
10.6229 | -104.857
10.6266 | -104.8456
10.6303 | -104.839
10.634 | -104.8386
10.6377 | -104.8381
10.6413 | -104.8539
10.645 | -104.8595
10.6487 | -104.8694
10.6524 | -104.8806
10.6561 | -104.8972
10.6598 | -104.9163
10.6635 | -104.9359
10.6672 | -104.9572
10.6708 | -104.978
10.6745 | -104.999
10.6782 | -105.0223
10.6819 | -105.0521
10.6856 | -105.0785
10.6893 | -105.1051
10.693 | -105.1314
10.6967 | -105.1544
10.7004 | -105.1728
10.704 | -105.186
10.7077 | -105.1978
10.7114 | -105.2029
10.7151 | -105.2071
10.7188 | -105.2107
10.7225 | -105.2109
10.7262 | -105.2042
10.7299 | -105.1969
10.7336 | -105.1851
10.7372 | -105.1716
10.7409 | -105.1593
10.7446 | -105.1437
10.7483 | -105.1325
10.752 | -105.1203
10.7557 | -105.1078
10.7594 | -105.1023
10.7631 | -105.095
10.7667 | -105.0886
10.7704 | -105.0863
10.7741 | -105.0864
10.7778 | -105.083
10.7815 | -105.0798
10.7852 | -105.0771
10.7889 | -105.0735
10.7926 | -105.0699
10.7963 | -105.07
10.7999 | -105.067
10.8036 | -105.0635
10.8073 | -105.0569
10.811 | -105.0466
10.8147 | -105.0351
10.8184 | -105.0195
10.8221 | -105.0042
10.8258 | -104.9879
10.8295 | -104.9734
10.8331 | -104.9554
10.8368 | -104.9354
10.8405 | -104.9144
10.8442 | -104.8955
10.8479 | -104.8788
10.8516 | -104.8675
10.8553 | -104.8599
10.859 | -104.8578
10.8626 | -104.8611
10.8663 | -104.8652
10.87 | -104.8714
10.8737 | -104.8778
10.8774 | -104.8884
10.8811 | -104.9018
10.8848 | -104.9164
10.8885 | -104.9322
10.8922 | -104.9504
10.8958 | -104.9679
10.8995 | -104.9822
10.9032 | -104.9973
10.9069 | -105.0064
10.9106 | -105.0111
10.9143 | -105.0167
10.918 | -105.0162
10.9217 | -105.0096
10.9254 | -105.0005
10.929 | -104.9884
10.9327 | -104.971
10.9364 | -104.9515
10.9401 | -104.9297
10.9438 | -104.9054
10.9475 | -104.8817
10.9512 | -104.8559
10.9549 | -104.8298
10.9586 | -104.8058
10.9622 | -104.7833
10.9659 | -104.7629
10.9696 | -104.7475
10.9733 | -104.7314
10.977 | -104.7218
10.9807 | -104.7158
10.9844 | -104.7124
10.9881 | -104.712
10.9917 | -104.7147
10.9954 | -104.7208
10.9991 | -104.7293
11.0028 | -104.7411
11.0065 | -104.7587
11.0102 | -104.7748
11.0139 | -104.7904
11.0176 | -104.8082
11.0213 | -104.8272
11.0249 | -104.8449
11.0286 | -104.864
11.0323 | -104.8838
11.036 | -104.8974
11.0397 | -104.91
11.0434 | -104.9243
11.0471 | -104.9331
11.0508 | -104.9397
11.0545 | -104.9448
11.0581 | -104.9477
11.0618 | -104.9494
11.0655 | -104.9505
11.0692 | -104.9476
11.0729 | -104.944
11.0766 | -104.9449
11.0803 | -104.9487
11.084 | -104.9493
11.0876 | -104.949
11.0913 | -104.9485
11.095 | -104.9514
11.0987 | -104.9545
11.1024 | -104.9562
11.1061 | -104.9562
11.1098 | -104.9525
11.1135 | -104.9456
11.1172 | -104.9363
11.1208 | -104.9266
11.1245 | -104.9226
11.1282 | -104.9139
11.1319 | -104.9056
11.1356 | -104.8994
11.1393 | -104.8931
11.143 | -104.887
11.1467 | -104.8832
11.1504 | -104.8842
11.154 | -104.8866
11.1577 | -104.8852
11.1614 | -104.8893
11.1651 | -104.8923
11.1688 | -104.8978
11.1725 | -104.9146
11.1762 | -104.9264
11.1799 | -104.9365
11.1835 | -104.9462
11.1872 | -104.9582
11.1909 | -104.9726
11.1946 | -104.9848
11.1983 | -104.9996
11.202 | -105.0137
11.2057 | -105.0295
11.2094 | -105.0444
11.2131 | -105.0592
11.2167 | -105.0713
11.2204 | -105.0873
11.2241 | -105.1056
11.2278 | -105.1236
11.2315 | -105.1457
11.2352 | -105.1642
11.2389 | -105.1769
11.2426 | -105.1902
11.2463 | -105.2057
11.2499 | -105.221
11.2536 | -105.2352
11.2573 | -105.2484
11.261 | -105.2599
11.2647 | -105.2703
11.2684 | -105.2803
11.2721 | -105.2916
11.2758 | -105.3008
11.2795 | -105.3127
11.2831 | -105.3205
11.2868 | -105.3219
11.2905 | -105.3237
11.2942 | -105.324
11.2979 | -105.3175
11.3016 | -105.3076
11.3053 | -105.2924
11.309 | -105.2723
11.3126 | -105.2463
11.3163 | -105.2176
11.32 | -105.189
11.3237 | -105.1537
11.3274 | -105.1129
11.3311 | -105.0694
11.3348 | -105.021
11.3385 | -104.9716
11.3422 | -104.9234
11.3458 | -104.8751
11.3495 | -104.8275
11.3532 | -104.7815
11.3569 | -104.736
11.3606 | -104.6955
11.3643 | -104.6597
11.368 | -104.6258
11.3717 | -104.5932
11.3754 | -104.5691
11.379 | -104.5481
11.3827 | -104.5351
11.3864 | -104.5244
11.3901 | -104.5172
11.3938 | -104.5147
11.3975 | -104.514
11.4012 | -104.5226
11.4049 | -104.5387
11.4085 | -104.5544
11.4122 | -104.572
11.4159 | -104.5941
11.4196 | -104.6211
11.4233 | -104.6465
11.427 | -104.6742
11.4307 | -104.7042
11.4344 | -104.7338
11.4381 | -104.7634
11.4417 | -104.7952
11.4454 | -104.8238
11.4491 | -104.8502
11.4528 | -104.876
11.4565 | -104.9017
11.4602 | -104.9209
11.4639 | -104.9393
11.4676 | -104.9616
11.4713 | -104.9819
11.4749 | -105.0008
11.4786 | -105.0179
11.4823 | -105.0293
11.486 | -105.0361
11.4897 | -105.0408
11.4934 | -105.0482
11.4971 | -105.0542
11.5008 | -105.0559
11.5044 | -105.0555
11.5081 | -105.0559
11.5118 | -105.0566
11.5155 | -105.0571
11.5192 | -105.0582
11.5229 | -105.0594
11.5266 | -105.059
11.5303 | -105.0586
11.534 | -105.0633
11.5376 | -105.0721
11.5413 | -105.0818
11.545 | -105.0925
11.5487 | -105.1022
11.5524 | -105.1135
11.5561 | -105.1283
11.5598 | -105.1463
11.5635 | -105.1704
11.5672 | -105.1904
11.5708 | -105.211
11.5745 | -105.2338
11.5782 | -105.2555
11.5819 | -105.2735
11.5856 | -105.2858
11.5893 | -105.3017
11.593 | -105.3129
11.5967 | -105.3183
11.6004 | -105.321
11.604 | -105.3231
11.6077 | -105.3225
11.6114 | -105.3166
11.6151 | -105.3106
11.6188 | -105.3059
11.6225 | -105.3007
11.6262 | -105.3013
11.6299 | -105.3017
11.6335 | -105.2989
11.6372 | -105.2975
11.6409 | -105.296
11.6446 | -105.2992
11.6483 | -105.3059
11.652 | -105.3129
11.6557 | -105.3169
11.6594 | -105.3216
11.6631 | -105.3296
11.6667 | -105.3399
11.6704 | -105.3485
11.6741 | -105.3543
11.6778 | -105.3578
11.6815 | -105.359
11.6852 | -105.3592
11.6889 | -105.3618
11.6926 | -105.3653
11.6963 | -105.3651
11.6999 | -105.3622
11.7036 | -105.3543
11.7073 | -105.3492
11.711 | -105.3438
11.7147 | -105.3387
11.7184 | -105.3335
11.7221 | -105.3333
11.7258 | -105.3331
11.7294 | -105.3304
11.7331 | -105.3254
11.7368 | -105.3261
11.7405 | -105.3294
11.7442 | -105.3361
11.7479 | -105.3428
11.7516 | -105.35
11.7553 | -105.3607
11.759 | -105.3698
11.7626 | -105.3784
11.7663 | -105.389
11.77 | -105.3994
11.7737 | -105.4101
11.7774 | -105.4179
11.7811 | -105.4297
11.7848 | -105.4404
11.7885 | -105.4449
11.7922 | -105.4486
11.7958 | -105.4485
11.7995 | -105.4431
11.8032 | -105.4361
11.8069 | -105.4312
11.8106 | -105.4243
11.8143 | -105.4094
11.818 | -105.3917
11.8217 | -105.3753
11.8253 | -105.3574
11.829 | -105.3353
11.8327 | -105.313
11.8364 | -105.2884
11.8401 | -105.2611
11.8438 | -105.2305
11.8475 | -105.2012
11.8512 | -105.1759
11.8549 | -105.1476
11.8585 | -105.118
11.8622 | -105.0931
11.8659 | -105.0702
11.8696 | -105.0513
11.8733 | -105.0338
11.877 | -105.0172
11.8807 | -105.002
11.8844 | -104.9895
11.8881 | -104.9805
11.8917 | -104.9729
11.8954 | -104.9662
11.8991 | -104.957
11.9028 | -104.9534
11.9065 | -104.949
11.9102 | -104.9439
11.9139 | -104.9387
11.9176 | -104.9318
11.9213 | -104.9273
11.9249 | -104.9219
11.9286 | -104.9144
11.9323 | -104.903
11.936 | -104.895
11.9397 | -104.8888
11.9434 | -104.8819
11.9471 | -104.8725
11.9508 | -104.8629
11.9544 | -104.8538
11.9581 | -104.8494
11.9618 | -104.8471
11.9655 | -104.8465
11.9692 | -104.8473
11.9729 | -104.8478
11.9766 | -104.853
11.9803 | -104.8574
11.984 | -104.8672
11.9876 | -104.8791
11.9913 | -104.8897
11.995 | -104.899
11.9987 | -104.9106
12.0024 | -104.9266
12.0061 | -104.9407
12.0098 | -104.9557
12.0135 | -104.9745
12.0172 | -104.9919
12.0208 | -105.0072
12.0245 | -105.02
12.0282 | -105.0339
12.0319 | -105.0484
12.0356 | -105.0576
12.0393 | -105.0649
12.043 | -105.0694
12.0467 | -105.0756
12.0503 | -105.0815
12.054 | -105.0884
12.0577 | -105.0916
12.0614 | -105.0918
12.0651 | -105.0936
12.0688 | -105.0925
12.0725 | -105.0937
12.0762 | -105.0908
12.0799 | -105.0913
12.0835 | -105.0911
12.0872 | -105.0893
12.0909 | -105.09
12.0946 | -105.09
12.0983 | -105.0892
12.102 | -105.0897
12.1057 | -105.088
12.1094 | -105.0836
12.1131 | -105.0778
12.1167 | -105.0681
12.1204 | -105.0627
12.1241 | -105.0616
12.1278 | -105.0589
12.1315 | -105.0507
12.1352 | -105.0419
12.1389 | -105.0351
12.1426 | -105.0275
12.1462 | -105.0221
12.1499 | -105.0113
12.1536 | -105.0038
12.1573 | -104.9926
12.161 | -104.9786
12.1647 | -104.9685
12.1684 | -104.9586
12.1721 | -104.9466
12.1758 | -104.9341
12.1794 | -104.9226
12.1831 | -104.9132
12.1868 | -104.9043
12.1905 | -104.8997
12.1942 | -104.898
12.1979 | -104.896
12.2016 | -104.8968
12.2053 | -104.9011
12.209 | -104.9129
12.2126 | -104.9286
12.2163 | -104.9433
12.22 | -104.9589
12.2237 | -104.9784
12.2274 | -104.9952
12.2311 | -105.018
12.2348 | -105.0447
12.2385 | -105.0724
12.2422 | -105.0995
12.2458 | -105.1241
12.2495 | -105.1492
12.2532 | -105.1792
12.2569 | -105.2064
12.2606 | -105.2319
12.2643 | -105.2604
12.268 | -105.2875
12.2717 | -105.3133
12.2753 | -105.3408
12.279 | -105.3676
12.2827 | -105.3984
12.2864 | -105.4274
12.2901 | -105.4545
12.2938 | -105.4804
12.2975 | -105.5006
12.3012 | -105.5151
12.3049 | -105.5279
12.3085 | -105.5428
12.3122 | -105.5541
12.3159 | -105.5574
12.3196 | -105.5608
12.3233 | -105.5643
12.327 | -105.5641
12.3307 | -105.5569
12.3344 | -105.5433
12.3381 | -105.5295
12.3417 | -105.5124
12.3454 | -105.4972
12.3491 | -105.4787
12.3528 | -105.456
12.3565 | -105.4364
12.3602 | -105.4151
12.3639 | -105.3899
12.3676 | -105.3652
12.3712 | -105.3409
12.3749 | -105.319
12.3786 | -105.2976
12.3823 | -105.2726
12.386 | -105.2468
12.3897 | -105.2239
12.3934 | -105.2043
12.3971 | -105.1862
12.4008 | -105.1686
12.4044 | -105.1565
12.4081 | -105.1437
12.4118 | -105.1331
12.4155 | -105.126
12.4192 | -105.1216
12.4229 | -105.1189
12.4266 | -105.1232
12.4303 | -105.1301
12.434 | -105.136
12.4376 | -105.1463
12.4413 | -105.1588
12.445 | -105.1719
12.4487 | -105.1837
12.4524 | -105.1958
12.4561 | -105.2079
12.4598 | -105.2179
12.4635 | -105.2287
12.4671 | -105.2363
12.4708 | -105.2429
12.4745 | -105.2536
12.4782 | -105.2627
12.4819 | -105.2745
12.4856 | -105.283
12.4893 | -105.2911
12.493 | -105.3001
12.4967 | -105.3075
12.5003 | -105.3125
12.504 | -105.3168
12.5077 | -105.3202
12.5114 | -105.3211
12.5151 | -105.3257
12.5188 | -105.3287
12.5225 | -105.3273
12.5262 | -105.3202
12.5299 | -105.3154
12.5335 | -105.3081
12.5372 | -105.3009
12.5409 | -105.2965
12.5446 | -105.2928
12.5483 | -105.2887
12.552 | -105.2842
12.5557 | -105.2783
12.5594 | -105.2747
12.5631 | -105.2735
12.5667 | -105.2721
12.5704 | -105.273
12.5741 | -105.2752
12.5778 | -105.277
12.5815 | -105.2793
12.5852 | -105.2813
12.5889 | -105.283
12.5926 | -105.2833
12.5962 | -105.2848
12.5999 | -105.2842
12.6036 | -105.2813
12.6073 | -105.278
12.611 | -105.2738
12.6147 | -105.2712
12.6184 | -105.2674
12.6221 | -105.2648
12.6258 | -105.2625
12.6294 | -105.2641
12.6331 | -105.2679
12.6368 | -105.2731
12.6405 | -105.2808
12.6442 | -105.287
12.6479 | -105.2972
12.6516 | -105.3096
12.6553 | -105.3209
12.659 | -105.3334
12.6626 | -105.3466
12.6663 | -105.3603
12.67 | -105.3759
12.6737 | -105.3961
12.6774 | -105.421
12.6811 | -105.4474
12.6848 | -105.4695
12.6885 | -105.4931
12.6921 | -105.5208
12.6958 | -105.5466
12.6995 | -105.5689
12.7032 | -105.5892
12.7069 | -105.6119
12.7106 | -105.6353
12.7143 | -105.657
12.718 | -105.6712
12.7217 | -105.6812
12.7253 | -105.6895
12.729 | -105.6944
12.7327 | -105.695
12.7364 | -105.693
12.7401 | -105.6866
12.7438 | -105.6774
12.7475 | -105.6643
12.7512 | -105.6485
12.7549 | -105.6321
12.7585 | -105.6125
12.7622 | -105.5918
12.7659 | -105.5674
12.7696 | -105.5441
12.7733 | -105.5208
12.777 | -105.499
12.7807 | -105.4804
12.7844 | -105.4654
12.788 | -105.4496
12.7917 | -105.4351
12.7954 | -105.4258
12.7991 | -105.417
12.8028 | -105.4128
12.8065 | -105.4093
12.8102 | -105.4058
12.8139 | -105.4057
12.8176 | -105.4072
12.8212 | -105.4067
12.8249 | -105.4077
12.8286 | -105.413
12.8323 | -105.4168
12.836 | -105.4214
12.8397 | -105.4268
12.8434 | -105.4337
12.8471 | -105.4383
12.8508 | -105.4438
12.8544 | -105.4475
12.8581 | -105.4496
12.8618 | -105.4522
12.8655 | -105.4514
12.8692 | -105.4509
12.8729 | -105.4486
12.8766 | -105.4433
12.8803 | -105.4371
12.884 | -105.433
12.8876 | -105.4275
12.8913 | -105.4147
12.895 | -105.4002
12.8987 | -105.3885
12.9024 | -105.379
12.9061 | -105.368
12.9098 | -105.3549
12.9135 | -105.3402
12.9171 | -105.3297
12.9208 | -105.3219
12.9245 | -105.3137
12.9282 | -105.3081
12.9319 | -105.3055
12.9356 | -105.302
12.9393 | -105.2984
12.943 | -105.2963
12.9467 | -105.2955
12.9503 | -105.2998
12.954 | -105.304
12.9577 | -105.3123
12.9614 | -105.3221
12.9651 | -105.3295
12.9688 | -105.3348
12.9725 | -105.3403
12.9762 | -105.3485
12.9799 | -105.3563
12.9835 | -105.3637
12.9872 | -105.3733
12.9909 | -105.3815
12.9946 | -105.3927
12.9983 | -105.4044
13.002 | -105.4103
13.0057 | -105.4188
13.0094 | -105.4305
13.013 | -105.4403
13.0167 | -105.4469
13.0204 | -105.4552
13.0241 | -105.4588
13.0278 | -105.462
13.0315 | -105.4646
13.0352 | -105.4688
13.0389 | -105.4692
13.0426 | -105.466
13.0462 | -105.4624
13.0499 | -105.4585
13.0536 | -105.4566
13.0573 | -105.4526
13.061 | -105.4493
13.0647 | -105.4485
13.0684 | -105.4458
13.0721 | -105.4423
13.0758 | -105.4405
13.0794 | -105.4419
13.0831 | -105.4433
13.0868 | -105.443
13.0905 | -105.4408
13.0942 | -105.4437
13.0979 | -105.4497
13.1016 | -105.4582
13.1053 | -105.4626
13.1089 | -105.465
13.1126 | -105.4691
13.1163 | -105.474
13.12 | -105.4794
13.1237 | -105.4849
13.1274 | -105.4896
13.1311 | -105.4946
13.1348 | -105.5001
13.1385 | -105.5015
13.1421 | -105.4998
13.1458 | -105.4945
13.1495 | -105.4896
13.1532 | -105.4863
13.1569 | -105.4796
13.1606 | -105.471
13.1643 | -105.4592
13.168 | -105.4474
13.1717 | -105.4332
13.1753 | -105.4189
13.179 | -105.4047
13.1827 | -105.3896
13.1864 | -105.3753
13.1901 | -105.3636
13.1938 | -105.3497
13.1975 | -105.3347
13.2012 | -105.3176
13.2049 | -105.3045
13.2085 | -105.2942
13.2122 | -105.2857
13.2159 | -105.2774
13.2196 | -105.2701
13.2233 | -105.2652
13.227 | -105.2678
13.2307 | -105.2723
13.2344 | -105.2739
13.238 | -105.2782
13.2417 | -105.2858
13.2454 | -105.2997
13.2491 | -105.3123
13.2528 | -105.3223
13.2565 | -105.337
13.2602 | -105.3523
13.2639 | -105.3695
13.2676 | -105.3881
13.2712 | -105.4046
13.2749 | -105.4238
13.2786 | -105.4458
13.2823 | -105.4611
13.286 | -105.4754
13.2897 | -105.4888
13.2934 | -105.5024
13.2971 | -105.5145
13.3008 | -105.5132
13.3044 | -105.5217
13.3081 | -105.5302
13.3118 | -105.5352
13.3155 | -105.538
13.3192 | -105.5403
13.3229 | -105.5383
13.3266 | -105.5321
13.3303 | -105.5275
13.3339 | -105.5198
13.3376 | -105.5106
13.3413 | -105.4999
13.345 | -105.4883
13.3487 | -105.4765
13.3524 | -105.4639
13.3561 | -105.4496
13.3598 | -105.4334
13.3635 | -105.4145
13.3671 | -105.3918
13.3708 | -105.368
13.3745 | -105.3461
13.3782 | -105.3278
13.3819 | -105.3051
13.3856 | -105.2824
13.3893 | -105.2656
13.393 | -105.2483
13.3967 | -105.2271
13.4003 | -105.2087
13.404 | -105.1928
13.4077 | -105.1742
13.4114 | -105.1574
13.4151 | -105.1441
13.4188 | -105.1299
13.4225 | -105.118
13.4262 | -105.1097
13.4298 | -105.1031
13.4335 | -105.1
13.4372 | -105.0966
13.4409 | -105.0992
13.4446 | -105.1064
13.4483 | -105.1141
13.452 | -105.1258
13.4557 | -105.1373
13.4594 | -105.1531
13.463 | -105.175
13.4667 | -105.1969
13.4704 | -105.2199
13.4741 | -105.2509
13.4778 | -105.283
13.4815 | -105.3155
13.4852 | -105.3493
13.4889 | -105.3843
13.4926 | -105.4218
13.4962 | -105.4564
13.4999 | -105.4906
13.5036 | -105.5206
13.5073 | -105.5504
13.511 | -105.5815
13.5147 | -105.61
13.5184 | -105.6384
13.5221 | -105.663
13.5258 | -105.6813
13.5294 | -105.7002
13.5331 | -105.7115
13.5368 | -105.7192
13.5405 | -105.7236
13.5442 | -105.7233
13.5479 | -105.7202
13.5516 | -105.7152
13.5553 | -105.7078
13.5589 | -105.6975
13.5626 | -105.6874
13.5663 | -105.6781
13.57 | -105.6632
13.5737 | -105.6455
13.5774 | -105.6343
13.5811 | -105.6215
13.5848 | -105.6035
13.5885 | -105.5916
13.5921 | -105.5823
13.5958 | -105.5684
13.5995 | -105.5529
13.6032 | -105.5439
13.6069 | -105.5356
13.6106 | -105.5226
13.6143 | -105.5124
13.618 | -105.5052
13.6217 | -105.4998
13.6253 | -105.4926
13.629 | -105.4837
13.6327 | -105.4763
13.6364 | -105.4682
13.6401 | -105.4613
13.6438 | -105.4549
13.6475 | -105.4506
13.6512 | -105.4479
13.6548 | -105.443
13.6585 | -105.4386
13.6622 | -105.4364
13.6659 | -105.4321
13.6696 | -105.4276
13.6733 | -105.4283
13.677 | -105.4286
13.6807 | -105.4301
13.6844 | -105.4334
13.688 | -105.4344
13.6917 | -105.437
13.6954 | -105.4378
13.6991 | -105.4362
13.7028 | -105.438
13.7065 | -105.439
13.7102 | -105.4421
13.7139 | -105.4429
13.7176 | -105.4435
13.7212 | -105.4445
13.7249 | -105.4456
13.7286 | -105.4503
13.7323 | -105.4511
13.736 | -105.4512
13.7397 | -105.4528
13.7434 | -105.4542
13.7471 | -105.4589
13.7507 | -105.4626
13.7544 | -105.4661
13.7581 | -105.4688
13.7618 | -105.4692
13.7655 | -105.4731
13.7692 | -105.4771
13.7729 | -105.4804
13.7766 | -105.4841
13.7803 | -105.4861
13.7839 | -105.4877
13.7876 | -105.4904
13.7913 | -105.4949
13.795 | -105.5017
13.7987 | -105.5055
13.8024 | -105.5077
13.8061 | -105.512
13.8098 | -105.5162
13.8135 | -105.5216
13.8171 | -105.5276
13.8208 | -105.5337
13.8245 | -105.5405
13.8282 | -105.5485
13.8319 | -105.5581
13.8356 | -105.5681
13.8393 | -105.5752
13.843 | -105.5803
13.8467 | -105.5816
13.8503 | -105.5856
13.854 | -105.5931
13.8577 | -105.6009
13.8614 | -105.6077
13.8651 | -105.6136
13.8688 | -105.6164
13.8725 | -105.6186
13.8762 | -105.6221
13.8798 | -105.6266
13.8835 | -105.6319
13.8872 | -105.6376
13.8909 | -105.6405
13.8946 | -105.6444
13.8983 | -105.648
13.902 | -105.6459
13.9057 | -105.6421
13.9094 | -105.6376
13.913 | -105.6332
13.9167 | -105.6277
13.9204 | -105.6212
13.9241 | -105.6152
13.9278 | -105.6083
13.9315 | -105.6021
13.9352 | -105.5925
13.9389 | -105.5833
13.9426 | -105.5717
13.9462 | -105.5569
13.9499 | -105.5422
13.9536 | -105.5295
13.9573 | -105.5123
13.961 | -105.4899
13.9647 | -105.4653
13.9684 | -105.4426
13.9721 | -105.4224
13.9757 | -105.4036
13.9794 | -105.3884
13.9831 | -105.3683
13.9868 | -105.3469
13.9905 | -105.3294
13.9942 | -105.3137
13.9979 | -105.2968
14.0016 | -105.2781
14.0053 | -105.2634
14.0089 | -105.2491
14.0126 | -105.2336
14.0163 | -105.2208
14.02 | -105.21
14.0237 | -105.2008
14.0274 | -105.1927
14.0311 | -105.1863
14.0348 | -105.1814
14.0385 | -105.1801
14.0421 | -105.1854
14.0458 | -105.1935
14.0495 | -105.2022
14.0532 | -105.2125
14.0569 | -105.2243
14.0606 | -105.2389
14.0643 | -105.2535
14.068 | -105.274
14.0716 | -105.2966
14.0753 | -105.3158
14.079 | -105.338
14.0827 | -105.3612
14.0864 | -105.387
14.0901 | -105.4143
14.0938 | -105.4402
14.0975 | -105.4678
14.1012 | -105.4913
14.1048 | -105.5138
14.1085 | -105.5334
14.1122 | -105.5514
14.1159 | -105.5698
14.1196 | -105.5853
14.1233 | -105.5972
14.127 | -105.6042
14.1307 | -105.6084
14.1344 | -105.6128
14.138 | -105.6164
14.1417 | -105.6153
14.1454 | -105.6147
14.1491 | -105.6134
14.1528 | -105.6105
14.1565 | -105.6046
14.1602 | -105.6013
14.1639 | -105.5985
14.1676 | -105.5941
14.1712 | -105.5875
14.1749 | -105.5787
14.1786 | -105.5691
14.1823 | -105.559
14.186 | -105.552
14.1897 | -105.547
14.1934 | -105.5396
14.1971 | -105.5325
14.2007 | -105.5307
14.2044 | -105.5336
14.2081 | -105.5351
14.2118 | -105.5372
14.2155 | -105.5415
14.2192 | -105.5489
14.2229 | -105.5585
14.2266 | -105.5687
14.2303 | -105.5787
14.2339 | -105.5888
14.2376 | -105.6006
14.2413 | -105.6098
14.245 | -105.6198
14.2487 | -105.6345
14.2524 | -105.65
14.2561 | -105.664
14.2598 | -105.677
14.2635 | -105.6892
14.2671 | -105.7051
14.2708 | -105.7176
14.2745 | -105.729
14.2782 | -105.7431
14.2819 | -105.7533
14.2856 | -105.7607
14.2893 | -105.766
14.293 | -105.7715
14.2966 | -105.7771
14.3003 | -105.7781
14.304 | -105.7833
14.3077 | -105.7858
14.3114 | -105.7852
14.3151 | -105.7827
14.3188 | -105.7796
14.3225 | -105.7782
14.3262 | -105.7738
14.3298 | -105.7678
14.3335 | -105.7609
14.3372 | -105.75
14.3409 | -105.7396
14.3446 | -105.7328
14.3483 | -105.7198
14.352 | -105.7036
14.3557 | -105.6898
14.3594 | -105.678
14.363 | -105.6622
14.3667 | -105.6419
14.3704 | -105.6234
14.3741 | -105.6027
14.3778 | -105.5774
14.3815 | -105.5547
14.3852 | -105.5364
14.3889 | -105.5176
14.3925 | -105.4998
14.3962 | -105.4806
14.3999 | -105.4625
14.4036 | -105.447
14.4073 | -105.4315
14.411 | -105.4172
14.4147 | -105.405
14.4184 | -105.3942
14.4221 | -105.3829
14.4257 | -105.3714
14.4294 | -105.3606
14.4331 | -105.355
14.4368 | -105.356
14.4405 | -105.3571
14.4442 | -105.3599
14.4479 | -105.3651
14.4516 | -105.3685
14.4553 | -105.3711
14.4589 | -105.3784
14.4626 | -105.3879
14.4663 | -105.3971
14.47 | -105.4044
14.4737 | -105.4127
14.4774 | -105.4263
14.4811 | -105.4367
14.4848 | -105.4467
14.4885 | -105.4592
14.4921 | -105.4673
14.4958 | -105.4742
14.4995 | -105.4803
14.5032 | -105.4883
14.5069 | -105.4947
14.5106 | -105.4973
14.5143 | -105.4975
14.518 | -105.496
14.5216 | -105.4925
14.5253 | -105.4875
14.529 | -105.4821
14.5327 | -105.4783
14.5364 | -105.473
14.5401 | -105.4653
14.5438 | -105.4577
14.5475 | -105.4474
14.5512 | -105.4321
14.5548 | -105.4201
14.5585 | -105.4117
14.5622 | -105.3984
14.5659 | -105.3833
14.5696 | -105.3716
14.5733 | -105.358
14.577 | -105.3478
14.5807 | -105.3394
14.5844 | -105.3306
14.588 | -105.3228
14.5917 | -105.3118
14.5954 | -105.3
14.5991 | -105.2895
14.6028 | -105.2806
14.6065 | -105.2717
14.6102 | -105.2672
14.6139 | -105.2638
14.6175 | -105.2601
14.6212 | -105.2579
14.6249 | -105.2558
14.6286 | -105.2576
14.6323 | -105.2619
14.636 | -105.2661
14.6397 | -105.2736
14.6434 | -105.2796
14.6471 | -105.2825
14.6507 | -105.2884
14.6544 | -105.2953
14.6581 | -105.3018
14.6618 | -105.3108
14.6655 | -105.3205
14.6692 | -105.3279
14.6729 | -105.3351
14.6766 | -105.345
14.6803 | -105.3592
14.6839 | -105.3723
14.6876 | -105.3803
14.6913 | -105.389
14.695 | -105.3974
14.6987 | -105.4082
14.7024 | -105.4252
14.7061 | -105.4414
14.7098 | -105.4548
14.7134 | -105.465
14.7171 | -105.4783
14.7208 | -105.4925
14.7245 | -105.5028
14.7282 | -105.5132
14.7319 | -105.5228
14.7356 | -105.531
14.7393 | -105.5376
14.743 | -105.5438
14.7466 | -105.5506
14.7503 | -105.5589
14.754 | -105.5657
14.7577 | -105.5718
14.7614 | -105.5793
14.7651 | -105.5831
14.7688 | -105.5839
14.7725 | -105.586
14.7762 | -105.5884
14.7798 | -105.5851
14.7835 | -105.5841
14.7872 | -105.5839
14.7909 | -105.5819
14.7946 | -105.5813
14.7983 | -105.5821
14.802 | -105.582
14.8057 | -105.5803
14.8094 | -105.5775
14.813 | -105.5757
14.8167 | -105.5752
14.8204 | -105.5721
14.8241 | -105.5684
14.8278 | -105.5658
14.8315 | -105.5624
14.8352 | -105.5614
14.8389 | -105.5593
14.8425 | -105.5565
14.8462 | -105.5561
14.8499 | -105.5552
14.8536 | -105.5548
14.8573 | -105.5531
14.861 | -105.5508
14.8647 | -105.5479
14.8684 | -105.5473
14.8721 | -105.5463
14.8757 | -105.546
14.8794 | -105.544
14.8831 | -105.5404
14.8868 | -105.5368
14.8905 | -105.5318
14.8942 | -105.5294
14.8979 | -105.5333
14.9016 | -105.5349
14.9053 | -105.539
14.9089 | -105.5433
14.9126 | -105.5435
14.9163 | -105.5428
14.92 | -105.5441
14.9237 | -105.55
14.9274 | -105.5589
14.9311 | -105.565
14.9348 | -105.5714
14.9384 | -105.5846
14.9421 | -105.5965
14.9458 | -105.6074
14.9495 | -105.6211
14.9532 | -105.6362
14.9569 | -105.6507
14.9606 | -105.6671
14.9643 | -105.6837
14.968 | -105.7006
14.9716 | -105.7168
14.9753 | -105.7339
14.979 | -105.7499
14.9827 | -105.764
14.9864 | -105.7768
14.9901 | -105.7879
14.9938 | -105.804
14.9975 | -105.8177
15.0012 | -105.8273
15.0048 | -105.8345
15.0085 | -105.8398
15.0122 | -105.8438
15.0159 | -105.8474
15.0196 | -105.8518
15.0233 | -105.8592
15.027 | -105.8639
15.0307 | -105.8655
15.0343 | -105.865
15.038 | -105.8616
15.0417 | -105.8598
15.0454 | -105.8575
15.0491 | -105.8536
15.0528 | -105.8496
15.0565 | -105.8454
15.0602 | -105.8383
15.0639 | -105.8271
15.0675 | -105.8174
15.0712 | -105.8107
15.0749 | -105.8054
15.0786 | -105.8003
15.0823 | -105.7912
15.086 | -105.7813
15.0897 | -105.7729
15.0934 | -105.7666
15.0971 | -105.7612
15.1007 | -105.7537
15.1044 | -105.7487
15.1081 | -105.7412
15.1118 | -105.7309
15.1155 | -105.7222
15.1192 | -105.7119
15.1229 | -105.703
15.1266 | -105.6956
15.1303 | -105.6881
15.1339 | -105.6792
15.1376 | -105.6705
15.1413 | -105.6616
15.145 | -105.652
15.1487 | -105.6442
15.1524 | -105.6376
15.1561 | -105.6282
15.1598 | -105.6212
15.1634 | -105.614
15.1671 | -105.6015
15.1708 | -105.5901
15.1745 | -105.5796
15.1782 | -105.5668
15.1819 | -105.5542
15.1856 | -105.545
15.1893 | -105.5396
15.193 | -105.5291
15.1966 | -105.5199
15.2003 | -105.5122
15.204 | -105.5044
15.2077 | -105.4977
15.2114 | -105.4896
15.2151 | -105.4812
15.2188 | -105.4756
15.2225 | -105.4708
15.2262 | -105.468
15.2298 | -105.465
15.2335 | -105.4624
15.2372 | -105.4598
15.2409 | -105.4568
15.2446 | -105.4558
15.2483 | -105.4565
15.252 | -105.458
15.2557 | -105.4625
15.2593 | -105.467
15.263 | -105.4717
15.2667 | -105.4756
15.2704 | -105.4813
15.2741 | -105.4866
15.2778 | -105.492
15.2815 | -105.5029
15.2852 | -105.5119
15.2889 | -105.5173
15.2925 | -105.5202
15.2962 | -105.5242
15.2999 | -105.5337
15.3036 | -105.5445
15.3073 | -105.5517
15.311 | -105.559
15.3147 | -105.567
15.3184 | -105.5748
15.3221 | -105.5821
15.3257 | -105.5916
15.3294 | -105.6039
15.3331 | -105.615
15.3368 | -105.6293
15.3405 | -105.6422
15.3442 | -105.6538
15.3479 | -105.6663
15.3516 | -105.6817
15.3552 | -105.6974
15.3589 | -105.7121
15.3626 | -105.7276
15.3663 | -105.744
15.37 | -105.7601
15.3737 | -105.7757
15.3774 | -105.7899
15.3811 | -105.8037
15.3848 | -105.82
15.3884 | -105.8333
15.3921 | -105.8459
15.3958 | -105.8607
15.3995 | -105.8741
15.4032 | -105.8825
15.4069 | -105.8891
15.4106 | -105.8959
15.4143 | -105.9007
15.418 | -105.9041
15.4216 | -105.9045
15.4253 | -105.9012
15.429 | -105.8798
15.4327 | -105.8728
15.4364 | -105.8639
15.4401 | -105.851
15.4438 | -105.8382
15.4475 | -105.8217
15.4512 | -105.8051
15.4548 | -105.7856
15.4585 | -105.7691
15.4622 | -105.7524
15.4659 | -105.7348
15.4696 | -105.7155
15.4733 | -105.6964
15.477 | -105.6789
15.4807 | -105.6607
15.4843 | -105.6422
15.488 | -105.6248
15.4917 | -105.6056
15.4954 | -105.5901
15.4991 | -105.5778
15.5028 | -105.5657
15.5065 | -105.5568
15.5102 | -105.5504
15.5139 | -105.5468
15.5175 | -105.5391
15.5212 | -105.5297
15.5249 | -105.5229
15.5286 | -105.521
15.5323 | -105.5212
15.536 | -105.5201
15.5397 | -105.5205
15.5434 | -105.5213
15.5471 | -105.5234
15.5507 | -105.5247
15.5544 | -105.5269
15.5581 | -105.529
15.5618 | -105.5331
15.5655 | -105.5387
15.5692 | -105.5443
15.5729 | -105.5497
15.5766 | -105.5542
15.5802 | -105.5577
15.5839 | -105.5613
15.5876 | -105.566
15.5913 | -105.5723
15.595 | -105.5777
15.5987 | -105.5815
15.6024 | -105.5855
15.6061 | -105.5895
15.6098 | -105.5957
15.6134 | -105.6006
15.6171 | -105.6018
15.6208 | -105.6056
15.6245 | -105.6136
15.6282 | -105.6197
15.6319 | -105.6276
15.6356 | -105.6363
15.6393 | -105.6441
15.643 | -105.6556
15.6466 | -105.6695
15.6503 | -105.6855
15.654 | -105.6984
15.6577 | -105.7111
15.6614 | -105.7255
15.6651 | -105.7415
15.6688 | -105.7554
15.6725 | -105.7685
15.6761 | -105.7801
15.6798 | -105.7936
15.6835 | -105.8026
15.6872 | -105.8094
15.6909 | -105.8168
15.6946 | -105.8253
15.6983 | -105.8321
15.702 | -105.8351
15.7057 | -105.8362
15.7093 | -105.8377
15.713 | -105.837
15.7167 | -105.8363
15.7204 | -105.832
15.7241 | -105.8254
15.7278 | -105.816
15.7315 | -105.8016
15.7352 | -105.7866
15.7389 | -105.7726
15.7425 | -105.7626
15.7462 | -105.7512
15.7499 | -105.7334
15.7536 | -105.7152
15.7573 | -105.6966
15.761 | -105.6777
15.7647 | -105.6589
15.7684 | -105.6407
15.7721 | -105.6214
15.7757 | -105.6053
15.7794 | -105.5887
15.7831 | -105.5715
15.7868 | -105.5515
15.7905 | -105.5325
15.7942 | -105.5159
15.7979 | -105.5
15.8016 | -105.4865
15.8052 | -105.4736
15.8089 | -105.4614
15.8126 | -105.4502
15.8163 | -105.4386
15.82 | -105.4277
15.8237 | -105.4211
15.8274 | -105.4162
15.8311 | -105.4097
15.8348 | -105.4035
15.8384 | -105.4008
15.8421 | -105.4004
15.8458 | -105.3974
15.8495 | -105.3944
15.8532 | -105.3944
15.8569 | -105.3978
15.8606 | -105.4026
15.8643 | -105.408
15.868 | -105.4123
15.8716 | -105.4151
15.8753 | -105.4207
15.879 | -105.4285
15.8827 | -105.4352
15.8864 | -105.4429
15.8901 | -105.4513
15.8938 | -105.4618
15.8975 | -105.4714
15.9011 | -105.4821
15.9048 | -105.4908
15.9085 | -105.4997
15.9122 | -105.509
15.9159 | -105.5168
15.9196 | -105.5243
15.9233 | -105.5317
15.927 | -105.5389
15.9307 | -105.5461
15.9343 | -105.5531
15.938 | -105.5606
15.9417 | -105.5668
15.9454 | -105.5739
15.9491 | -105.5808
15.9528 | -105.5885
15.9565 | -105.5995
15.9602 | -105.6121
15.9639 | -105.6223
15.9675 | -105.6319
15.9712 | -105.6379
15.9749 | -105.6453
15.9786 | -105.6564
15.9823 | -105.6652
15.986 | -105.6736
15.9897 | -105.6841
15.9934 | -105.6927
15.997 | -105.6998
16.0007 | -105.7074
16.0044 | -105.7142
16.0081 | -105.7209
16.0118 | -105.729
16.0155 | -105.7392
16.0192 | -105.7487
16.0229 | -105.7564
16.0266 | -105.7618
16.0302 | -105.7688
16.0339 | -105.7766
16.0376 | -105.7812
16.0413 | -105.7863
16.045 | -105.7941
16.0487 | -105.8008
16.0524 | -105.8051
16.0561 | -105.8089
16.0598 | -105.8112
16.0634 | -105.8148
16.0671 | -105.8161
16.0708 | -105.8153
16.0745 | -105.8152
16.0782 | -105.8093
16.0819 | -105.8044
16.0856 | -105.7995
16.0893 | -105.7961
16.0929 | -105.7926
16.0966 | -105.789
16.1003 | -105.7812
16.104 | -105.7718
16.1077 | -105.7612
16.1114 | -105.7507
16.1151 | -105.742
16.1188 | -105.734
16.1225 | -105.7229
16.1261 | -105.7144
16.1298 | -105.7035
16.1335 | -105.6931
16.1372 | -105.6833
16.1409 | -105.6712
16.1446 | -105.6589
16.1483 | -105.648
16.152 | -105.6365
16.1557 | -105.6242
16.1593 | -105.6133
16.163 | -105.6002
16.1667 | -105.5873
16.1704 | -105.5775
16.1741 | -105.5669
16.1778 | -105.5555
16.1815 | -105.5454
16.1852 | -105.5335
16.1889 | -105.5216
16.1925 | -105.5121
16.1962 | -105.5006
16.1999 | -105.4879
16.2036 | -105.4802
16.2073 | -105.4742
16.211 | -105.467
16.2147 | -105.4607
16.2184 | -105.455
16.222 | -105.4485
16.2257 | -105.4442
16.2294 | -105.4391
16.2331 | -105.4326
16.2368 | -105.428
16.2405 | -105.4259
16.2442 | -105.4239
16.2479 | -105.4208
16.2516 | -105.4163
16.2552 | -105.4141
16.2589 | -105.4155
16.2626 | -105.416
16.2663 | -105.4154
16.27 | -105.4139
16.2737 | -105.4117
16.2774 | -105.4104
16.2811 | -105.4099
16.2848 | -105.4109
16.2884 | -105.4124
16.2921 | -105.4165
16.2958 | -105.4194
16.2995 | -105.4222
16.3032 | -105.4254
16.3069 | -105.4282
16.3106 | -105.4358
16.3143 | -105.4412
16.3179 | -105.4457
16.3216 | -105.4481
16.3253 | -105.4514
16.329 | -105.4553
16.3327 | -105.4657
16.3364 | -105.476
16.3401 | -105.4826
16.3438 | -105.4886
16.3475 | -105.4949
16.3511 | -105.4991
16.3548 | -105.5038
16.3585 | -105.5081
16.3622 | -105.5152
16.3659 | -105.5216
16.3696 | -105.5246
16.3733 | -105.5292
16.377 | -105.5323
16.3807 | -105.5347
16.3843 | -105.5374
16.388 | -105.5393
16.3917 | -105.5396
16.3954 | -105.5413
16.3991 | -105.5386
16.4028 | -105.5342
16.4065 | -105.5324
16.4102 | -105.5313
16.4139 | -105.5277
16.4175 | -105.5242
16.4212 | -105.5218
16.4249 | -105.5203
16.4286 | -105.5208
16.4323 | -105.5209
16.436 | -105.5186
16.4397 | -105.5158
16.4434 | -105.5155
16.447 | -105.5159
16.4507 | -105.5181
16.4544 | -105.5208
16.4581 | -105.5259
16.4618 | -105.5337
16.4655 | -105.5398
16.4692 | -105.5467
16.4729 | -105.5539
16.4766 | -105.562
16.4802 | -105.5736
16.4839 | -105.5871
16.4876 | -105.596
16.4913 | -105.6022
16.495 | -105.6131
16.4987 | -105.6212
16.5024 | -105.6309
16.5061 | -105.6437
16.5098 | -105.654
16.5134 | -105.6625
16.5171 | -105.672
16.5208 | -105.6781
16.5245 | -105.6853
16.5282 | -105.6902
16.5319 | -105.694
16.5356 | -105.6969
16.5393 | -105.6993
16.5429 | -105.7012
16.5466 | -105.7
16.5503 | -105.6972
16.554 | -105.6926
16.5577 | -105.6883
16.5614 | -105.6863
16.5651 | -105.6847
16.5688 | -105.6795
16.5725 | -105.6741
16.5761 | -105.6721
16.5798 | -105.6675
16.5835 | -105.6631
16.5872 | -105.6614
16.5909 | -105.6613
16.5946 | -105.6597
16.5983 | -105.6556
16.602 | -105.6572
16.6057 | -105.6607
16.6093 | -105.6639
16.613 | -105.6675
16.6167 | -105.6737
16.6204 | -105.6804
16.6241 | -105.6878
16.6278 | -105.6967
16.6315 | -105.7053
16.6352 | -105.7177
16.6388 | -105.7298
16.6425 | -105.7396
16.6462 | -105.7503
16.6499 | -105.7623
16.6536 | -105.7754
16.6573 | -105.7888
16.661 | -105.8005
16.6647 | -105.8108
16.6684 | -105.8224
16.672 | -105.8335
16.6757 | -105.8443
16.6794 | -105.8557
16.6831 | -105.8666
16.6868 | -105.8773
16.6905 | -105.8877
16.6942 | -105.8959
16.6979 | -105.9034
16.7016 | -105.9166
16.7052 | -105.9271
16.7089 | -105.9344
16.7126 | -105.9406
16.7163 | -105.9455
16.72 | -105.9515
16.7237 | -105.9602
16.7274 | -105.9697
16.7311 | -105.973
16.7347 | -105.9739
16.7384 | -105.9756
16.7421 | -105.9778
16.7458 | -105.9736
16.7495 | -105.9693
16.7532 | -105.9648
16.7569 | -105.9606
16.7606 | -105.9567
16.7643 | -105.949
16.7679 | -105.9439
16.7716 | -105.9365
16.7753 | -105.9306
16.779 | -105.9242
16.7827 | -105.9164
16.7864 | -105.9107
16.7901 | -105.9051
16.7938 | -105.9001
16.7975 | -105.8962
16.8011 | -105.8884
16.8048 | -105.8802
16.8085 | -105.8743
16.8122 | -105.8687
16.8159 | -105.8657
16.8196 | -105.8649
16.8233 | -105.8612
16.827 | -105.8581
16.8307 | -105.8564
16.8343 | -105.8572
16.838 | -105.86
16.8417 | -105.8581
16.8454 | -105.8566
16.8491 | -105.8598
16.8528 | -105.8614
16.8565 | -105.8582
16.8602 | -105.857
16.8638 | -105.8596
16.8675 | -105.8627
16.8712 | -105.8666
16.8749 | -105.8686
16.8786 | -105.8711
16.8823 | -105.8755
16.886 | -105.8818
16.8897 | -105.8891
16.8934 | -105.8975
16.897 | -105.9027
16.9007 | -105.9099
16.9044 | -105.9189
16.9081 | -105.9273
16.9118 | -105.9356
16.9155 | -105.9456
16.9192 | -105.9528
16.9229 | -105.9608
16.9266 | -105.9693
16.9302 | -105.978
16.9339 | -105.9859
16.9376 | -105.9931
16.9413 | -105.9983
16.945 | -106.0045
16.9487 | -106.0138
16.9524 | -106.0224
16.9561 | -106.0285
16.9597 | -106.0317
16.9634 | -106.032
16.9671 | -106.0289
16.9708 | -106.0274
16.9745 | -106.0266
16.9782 | -106.0245
16.9819 | -106.0198
16.9856 | -106.0169
16.9893 | -106.0122
16.9929 | -106.0051
16.9966 | -105.9959
17.0003 | -105.9877
17.004 | -105.9805
17.0077 | -105.9678
17.0114 | -105.9579
17.0151 | -105.9476
17.0188 | -105.9355
17.0225 | -105.9225
17.0261 | -105.9068
17.0298 | -105.8937
17.0335 | -105.8784
17.0372 | -105.8629
17.0409 | -105.8472
17.0446 | -105.8339
17.0483 | -105.8182
17.052 | -105.805
17.0556 | -105.793
17.0593 | -105.7812
17.063 | -105.7696
17.0667 | -105.7596
17.0704 | -105.7505
17.0741 | -105.7419
17.0778 | -105.7318
17.0815 | -105.7227
17.0852 | -105.7167
17.0888 | -105.7126
17.0925 | -105.707
17.0962 | -105.706
17.0999 | -105.7087
17.1036 | -105.7092
17.1073 | -105.71
17.111 | -105.7136
17.1147 | -105.7189
17.1184 | -105.7262
17.122 | -105.7353
17.1257 | -105.7447
17.1294 | -105.7547
17.1331 | -105.768
17.1368 | -105.7835
17.1405 | -105.7997
17.1442 | -105.8168
17.1479 | -105.8353
17.1516 | -105.8562
17.1552 | -105.8773
17.1589 | -105.8975
17.1626 | -105.9187
17.1663 | -105.9403
17.17 | -105.9616
17.1737 | -105.9843
17.1774 | -106.0074
17.1811 | -106.0302
17.1847 | -106.0507
17.1884 | -106.0709
17.1921 | -106.0895
17.1958 | -106.1074
17.1995 | -106.1262
17.2032 | -106.1455
17.2069 | -106.1664
17.2106 | -106.1858
17.2143 | -106.2004
17.2179 | -106.2109
17.2216 | -106.2211
17.2253 | -106.2317
17.229 | -106.2412
17.2327 | -106.2478
17.2364 | -106.2517
17.2401 | -106.2558
17.2438 | -106.2553
17.2475 | -106.2521
17.2511 | -106.2511
17.2548 | -106.2459
17.2585 | -106.2351
17.2622 | -106.2261
17.2659 | -106.2176
17.2696 | -106.2081
17.2733 | -106.1974
17.277 | -106.1854
17.2806 | -106.1742
17.2843 | -106.162
17.288 | -106.1484
17.2917 | -106.1345
17.2954 | -106.1245
17.2991 | -106.1118
17.3028 | -106.0969
17.3065 | -106.0822
17.3102 | -106.0675
17.3138 | -106.0521
17.3175 | -106.0376
17.3212 | -106.0251
17.3249 | -106.0158
17.3286 | -106.0057
17.3323 | -105.9947
17.336 | -105.9859
17.3397 | -105.9781
17.3434 | -105.9694
17.347 | -105.9611
17.3507 | -105.9558
17.3544 | -105.9513
17.3581 | -105.9466
17.3618 | -105.943
17.3655 | -105.9393
17.3692 | -105.9335
17.3729 | -105.9305
17.3766 | -105.9257
17.3802 | -105.9216
17.3839 | -105.9194
17.3876 | -105.9203
17.3913 | -105.9206
17.395 | -105.9208
17.3987 | -105.921
17.4024 | -105.921
17.4061 | -105.9193
17.4097 | -105.9204
17.4134 | -105.9213
17.4171 | -105.9221
17.4208 | -105.9235
17.4245 | -105.9249
17.4282 | -105.9259
17.4319 | -105.927
17.4356 | -105.9283
17.4393 | -105.93
17.4429 | -105.9329
17.4466 | -105.9366
17.4503 | -105.9385
17.454 | -105.9391
17.4577 | -105.9421
17.4614 | -105.9482
17.4651 | -105.954
17.4688 | -105.957
17.4725 | -105.9585
17.4761 | -105.9618
17.4798 | -105.9629
17.4835 | -105.9665
17.4872 | -105.9709
17.4909 | -105.9771
17.4946 | -105.9849
17.4983 | -105.9915
17.502 | -105.9966
17.5056 | -106.0055
17.5093 | -106.0112
17.513 | -106.0158
17.5167 | -106.0161
17.5204 | -106.0153
17.5241 | -106.0152
17.5278 | -106.0197
17.5315 | -106.0252
17.5352 | -106.0308
17.5388 | -106.0325
17.5425 | -106.0326
17.5462 | -106.033
17.5499 | -106.0335
17.5536 | -106.0297
17.5573 | -106.0182
17.561 | -106.0154
17.5647 | -106.0095
17.5684 | -106.0033
17.572 | -105.9965
17.5757 | -105.9871
17.5794 | -105.9781
17.5831 | -105.9698
17.5868 | -105.9576
17.5905 | -105.9438
17.5942 | -105.9302
17.5979 | -105.9126
17.6015 | -105.8975
17.6052 | -105.8817
17.6089 | -105.8623
17.6126 | -105.8435
17.6163 | -105.83
17.62 | -105.8169
17.6237 | -105.801
17.6274 | -105.7825
17.6311 | -105.7635
17.6347 | -105.7433
17.6384 | -105.7239
17.6421 | -105.7048
17.6458 | -105.6888
17.6495 | -105.6728
17.6532 | -105.6573
17.6569 | -105.6436
17.6606 | -105.6294
17.6643 | -105.6168
17.6679 | -105.6048
17.6716 | -105.5919
17.6753 | -105.5798
17.679 | -105.5681
17.6827 | -105.5573
17.6864 | -105.5474
17.6901 | -105.5384
17.6938 | -105.5289
17.6974 | -105.5184
17.7011 | -105.5103
17.7048 | -105.5029
17.7085 | -105.495
17.7122 | -105.4863
17.7159 | -105.4796
17.7196 | -105.4747
17.7233 | -105.4721
17.727 | -105.4681
17.7306 | -105.4665
17.7343 | -105.464
17.738 | -105.4607
17.7417 | -105.4584
17.7454 | -105.4554
17.7491 | -105.4552
17.7528 | -105.4587
17.7565 | -105.4612
17.7602 | -105.4656
17.7638 | -105.4704
17.7675 | -105.4782
17.7712 | -105.4861
17.7749 | -105.4939
17.7786 | -105.4996
17.7823 | -105.5057
17.786 | -105.5144
17.7897 | -105.5276
17.7934 | -105.5405
17.797 | -105.5526
17.8007 | -105.5669
17.8044 | -105.5847
17.8081 | -105.6017
17.8118 | -105.6165
17.8155 | -105.6332
17.8192 | -105.6525
17.8229 | -105.6687
17.8265 | -105.6835
17.8302 | -105.7012
17.8339 | -105.7182
17.8376 | -105.7326
17.8413 | -105.7496
17.845 | -105.7688
17.8487 | -105.7877
17.8524 | -105.8096
17.8561 | -105.8268
17.8597 | -105.8445
17.8634 | -105.8662
17.8671 | -105.8875
17.8708 | -105.9054
17.8745 | -105.9243
17.8782 | -105.9417
17.8819 | -105.9617
17.8856 | -105.9787
17.8893 | -105.9948
17.8929 | -106.0094
17.8966 | -106.0251
17.9003 | -106.0393
17.904 | -106.0539
17.9077 | -106.0662
17.9114 | -106.0771
17.9151 | -106.085
17.9188 | -106.0927
17.9224 | -106.101
17.9261 | -106.1096
17.9298 | -106.1174
17.9335 | -106.1215
17.9372 | -106.1252
17.9409 | -106.1264
17.9446 | -106.1268
17.9483 | -106.1265
17.952 | -106.1242
17.9556 | -106.1206
17.9593 | -106.1179
17.963 | -106.1134
17.9667 | -106.1092
17.9704 | -106.104
17.9741 | -106.0974
17.9778 | -106.0856
17.9815 | -106.0746
17.9852 | -106.0642
17.9888 | -106.0536
17.9925 | -106.043
17.9962 | -106.0297
17.9999 | -106.016
18.0036 | -105.9993
18.0073 | -105.9819
18.011 | -105.9678
18.0147 | -105.9518
18.0184 | -105.9355
18.022 | -105.9185
18.0257 | -105.9019
18.0294 | -105.8853
18.0331 | -105.8692
18.0368 | -105.8515
18.0405 | -105.8307
18.0442 | -105.807
18.0479 | -105.7864
18.0515 | -105.7671
18.0552 | -105.7463
18.0589 | -105.7244
18.0626 | -105.7053
18.0663 | -105.6868
18.07 | -105.6673
18.0737 | -105.6471
18.0774 | -105.6274
18.0811 | -105.6074
18.0847 | -105.5856
18.0884 | -105.566
18.0921 | -105.5503
18.0958 | -105.5347
18.0995 | -105.5206
18.1032 | -105.5033
18.1069 | -105.4845
18.1106 | -105.4682
18.1143 | -105.4541
18.1179 | -105.4413
18.1216 | -105.4297
18.1253 | -105.4178
18.129 | -105.4077
18.1327 | -105.3995
18.1364 | -105.3962
18.1401 | -105.3954
18.1438 | -105.3959
18.1474 | -105.398
18.1511 | -105.4003
18.1548 | -105.4035
18.1585 | -105.4074
18.1622 | -105.4117
18.1659 | -105.4202
18.1696 | -105.4299
18.1733 | -105.4418
18.177 | -105.4505
18.1806 | -105.4614
18.1843 | -105.4743
18.188 | -105.4873
18.1917 | -105.5042
18.1954 | -105.5192
18.1991 | -105.5353
18.2028 | -105.5523
18.2065 | -105.5672
18.2102 | -105.5856
18.2138 | -105.6039
18.2175 | -105.6214
18.2212 | -105.6391
18.2249 | -105.6565
18.2286 | -105.6753
18.2323 | -105.6932
18.236 | -105.7116
18.2397 | -105.73
18.2433 | -105.747
18.247 | -105.7614
18.2507 | -105.7758
18.2544 | -105.7914
18.2581 | -105.8095
18.2618 | -105.8283
18.2655 | -105.8438
18.2692 | -105.854
18.2729 | -105.8641
18.2765 | -105.8729
18.2802 | -105.8817
18.2839 | -105.889
18.2876 | -105.8953
18.2913 | -105.9005
18.295 | -105.9069
18.2987 | -105.9132
18.3024 | -105.9208
18.3061 | -105.9263
18.3097 | -105.9293
18.3134 | -105.9306
18.3171 | -105.9327
18.3208 | -105.9319
18.3245 | -105.9293
18.3282 | -105.9268
18.3319 | -105.9283
18.3356 | -105.9276
18.3392 | -105.9237
18.3429 | -105.9204
18.3466 | -105.9163
18.3503 | -105.9108
18.354 | -105.9066
18.3577 | -105.903
18.3614 | -105.8998
18.3651 | -105.8991
18.3688 | -105.8982
18.3724 | -105.8951
18.3761 | -105.8934
18.3798 | -105.889
18.3835 | -105.8835
18.3872 | -105.8807
18.3909 | -105.8766
18.3946 | -105.8691
18.3983 | -105.8596
18.402 | -105.8512
18.4056 | -105.8429
18.4093 | -105.8342
18.413 | -105.8264
18.4167 | -105.8141
18.4204 | -105.7997
18.4241 | -105.7869
18.4278 | -105.7741
18.4315 | -105.7633
18.4352 | -105.7515
18.4388 | -105.7389
18.4425 | -105.7232
18.4462 | -105.7065
18.4499 | -105.6924
18.4536 | -105.6777
18.4573 | -105.6636
18.461 | -105.6464
18.4647 | -105.6312
18.4683 | -105.6193
18.472 | -105.6033
18.4757 | -105.5869
18.4794 | -105.5711
18.4831 | -105.5564
18.4868 | -105.5418
18.4905 | -105.5253
18.4942 | -105.5102
18.4979 | -105.4963
18.5015 | -105.4819
18.5052 | -105.4691
18.5089 | -105.4568
18.5126 | -105.4432
18.5163 | -105.4323
18.52 | -105.4214
18.5237 | -105.4093
18.5274 | -105.3996
18.5311 | -105.3925
18.5347 | -105.3882
18.5384 | -105.3855
18.5421 | -105.3817
18.5458 | -105.3772
18.5495 | -105.3745
18.5532 | -105.3751
18.5569 | -105.3754
18.5606 | -105.3769
18.5642 | -105.3782
18.5679 | -105.3788
18.5716 | -105.382
18.5753 | -105.3872
18.579 | -105.3915
18.5827 | -105.3985
18.5864 | -105.4037
18.5901 | -105.4082
18.5938 | -105.4119
18.5974 | -105.4185
18.6011 | -105.4251
18.6048 | -105.4319
18.6085 | -105.4369
18.6122 | -105.4418
18.6159 | -105.4461
18.6196 | -105.4471
18.6233 | -105.4449
18.627 | -105.4425
18.6306 | -105.4418
18.6343 | -105.4434
18.638 | -105.4457
18.6417 | -105.4452
18.6454 | -105.4447
18.6491 | -105.4442
18.6528 | -105.4411
18.6565 | -105.4374
18.6602 | -105.4369
18.6638 | -105.4364
18.6675 | -105.4336
18.6712 | -105.4296
18.6749 | -105.4228
18.6786 | -105.4175
18.6823 | -105.4151
18.686 | -105.413
18.6897 | -105.4071
18.6933 | -105.4016
18.697 | -105.3938
18.7007 | -105.3872
18.7044 | -105.3838
18.7081 | -105.3786
18.7118 | -105.3718
18.7155 | -105.3626
18.7192 | -105.3536
18.7229 | -105.3468
18.7265 | -105.3429
18.7302 | -105.3361
18.7339 | -105.3261
18.7376 | -105.3197
18.7413 | -105.3149
18.745 | -105.3083
18.7487 | -105.3002
18.7524 | -105.293
18.7561 | -105.2871
18.7597 | -105.2823
18.7634 | -105.2789
18.7671 | -105.2766
18.7708 | -105.2723
18.7745 | -105.2676
18.7782 | -105.2643
18.7819 | -105.2566
18.7856 | -105.2479
18.7892 | -105.2408
18.7929 | -105.2331
18.7966 | -105.2261
18.8003 | -105.2168
18.804 | -105.206
18.8077 | -105.1969
18.8114 | -105.1878
18.8151 | -105.1761
18.8188 | -105.1636
18.8224 | -105.1492
18.8261 | -105.1368
18.8298 | -105.1259
18.8335 | -105.1154
18.8372 | -105.1048
18.8409 | -105.0931
18.8446 | -105.0792
18.8483 | -105.0651
18.852 | -105.0554
18.8556 | -105.0475
18.8593 | -105.0355
18.863 | -105.0248
18.8667 | -105.0171
18.8704 | -105.0064
18.8741 | -104.9963
18.8778 | -104.9899
18.8815 | -104.9841
18.8851 | -104.9802
18.8888 | -104.9747
18.8925 | -104.9701
18.8962 | -104.9674
18.8999 | -104.9638
18.9036 | -104.9592
18.9073 | -104.955
18.911 | -104.9539
18.9147 | -104.9538
18.9183 | -104.9524
18.922 | -104.9527
18.9257 | -104.9492
18.9294 | -104.9434
18.9331 | -104.9371
18.9368 | -104.9326
18.9405 | -104.9284
18.9442 | -104.9253
18.9479 | -104.9189
18.9515 | -104.9082
18.9552 | -104.8988
18.9589 | -104.8926
18.9626 | -104.8827
18.9663 | -104.8704
18.97 | -104.8574
18.9737 | -104.845
18.9774 | -104.8328
18.981 | -104.8195
18.9847 | -104.8053
18.9884 | -104.7898
18.9921 | -104.7712
18.9958 | -104.7547
18.9995 | -104.7373
19.0032 | -104.7203
19.0069 | -104.7022
19.0106 | -104.6813
19.0142 | -104.6626
19.0179 | -104.6461
19.0216 | -104.6286
19.0253 | -104.6118
19.029 | -104.5957
19.0327 | -104.5786
19.0364 | -104.5618
19.0401 | -104.5478
19.0438 | -104.5359
19.0474 | -104.522
19.0511 | -104.5103
19.0548 | -104.4969
19.0585 | -104.4854
19.0622 | -104.4754
19.0659 | -104.4654
19.0696 | -104.4578
19.0733 | -104.4506
19.077 | -104.4423
19.0806 | -104.4366
19.0843 | -104.4325
19.088 | -104.4271
19.0917 | -104.4255
19.0954 | -104.4225
19.0991 | -104.4184
19.1028 | -104.4134
19.1065 | -104.4115
19.1101 | -104.4109
19.1138 | -104.4128
19.1175 | -104.4145
19.1212 | -104.4152
19.1249 | -104.414
19.1286 | -104.4169
19.1323 | -104.4206
19.136 | -104.4215
19.1397 | -104.423
19.1433 | -104.4262
19.147 | -104.4281
19.1507 | -104.4294
19.1544 | -104.4323
19.1581 | -104.4349
19.1618 | -104.4391
19.1655 | -104.4411
19.1692 | -104.4419
19.1729 | -104.4424
19.1765 | -104.4458
19.1802 | -104.447
19.1839 | -104.4466
19.1876 | -104.4438
19.1913 | -104.4407
19.195 | -104.4378
19.1987 | -104.4376
19.2024 | -104.4354
19.206 | -104.433
19.2097 | -104.4282
19.2134 | -104.4239
19.2171 | -104.4181
19.2208 | -104.4137
19.2245 | -104.4081
19.2282 | -104.4013
19.2319 | -104.3933
19.2356 | -104.3849
19.2392 | -104.3755
19.2429 | -104.3667
19.2466 | -104.357
19.2503 | -104.3502
19.254 | -104.3429
19.2577 | -104.3349
19.2614 | -104.325
19.2651 | -104.3148
19.2688 | -104.3034
19.2724 | -104.292
19.2761 | -104.2838
19.2798 | -104.276
19.2835 | -104.2681
19.2872 | -104.2581
19.2909 | -104.2487
19.2946 | -104.2392
19.2983 | -104.2319
19.302 | -104.2239
19.3056 | -104.2163
19.3093 | -104.211
19.313 | -104.206
19.3167 | -104.2022
19.3204 | -104.2022
19.3241 | -104.2029
19.3278 | -104.2066
19.3315 | -104.2107
19.3351 | -104.2136
19.3388 | -104.2185
19.3425 | -104.2235
19.3462 | -104.2288
19.3499 | -104.2352
19.3536 | -104.2407
19.3573 | -104.2473
19.361 | -104.2559
19.3647 | -104.2646
19.3683 | -104.2748
19.372 | -104.2852
19.3757 | -104.297
19.3794 | -104.3054
19.3831 | -104.3128
19.3868 | -104.3206
19.3905 | -104.33
19.3942 | -104.3377
19.3979 | -104.348
19.4015 | -104.3566
19.4052 | -104.3645
19.4089 | -104.373
19.4126 | -104.3792
19.4163 | -104.3836
19.42 | -104.3871
19.4237 | -104.3927
19.4274 | -104.3981
19.431 | -104.4027
19.4347 | -104.4077
19.4384 | -104.4124
19.4421 | -104.4183
19.4458 | -104.4238
19.4495 | -104.4284
19.4532 | -104.431
19.4569 | -104.4341
19.4606 | -104.4382
19.4642 | -104.4458
19.4679 | -104.4529
19.4716 | -104.4575
19.4753 | -104.4632
19.479 | -104.4683
19.4827 | -104.4721
19.4864 | -104.4796
19.4901 | -104.4864
19.4938 | -104.4936
19.4974 | -104.5003
19.5011 | -104.5091
19.5048 | -104.5173
19.5085 | -104.5255
19.5122 | -104.5329
19.5159 | -104.5412
19.5196 | -104.5502
19.5233 | -104.5605
19.5269 | -104.5681
19.5306 | -104.5757
19.5343 | -104.5854
19.538 | -104.5962
19.5417 | -104.6083
19.5454 | -104.6179
19.5491 | -104.6297
19.5528 | -104.639
19.5565 | -104.6475
19.5601 | -104.6589
19.5638 | -104.6712
19.5675 | -104.6814
19.5712 | -104.6926
19.5749 | -104.7009
19.5786 | -104.71
19.5823 | -104.7204
19.586 | -104.7284
19.5897 | -104.7372
19.5933 | -104.7457
19.597 | -104.7517
19.6007 | -104.7569
19.6044 | -104.7626
19.6081 | -104.7678
19.6118 | -104.772
19.6155 | -104.7744
19.6192 | -104.7787
19.6228 | -104.7855
19.6265 | -104.7922
19.6302 | -104.7973
19.6339 | -104.7982
19.6376 | -104.7986
19.6413 | -104.8014
19.645 | -104.8036
19.6487 | -104.805
19.6524 | -104.8049
19.656 | -104.8052
19.6597 | -104.8066
19.6634 | -104.8084
19.6671 | -104.8079
19.6708 | -104.8094
19.6745 | -104.8088
19.6782 | -104.8069
19.6819 | -104.805
19.6856 | -104.8036
19.6892 | -104.8044
19.6929 | -104.8058
19.6966 | -104.8062
19.7003 | -104.8079
19.704 | -104.8077
19.7077 | -104.8059
19.7114 | -104.805
19.7151 | -104.8051
19.7188 | -104.8056
19.7224 | -104.8091
19.7261 | -104.8118
19.7298 | -104.8156
19.7335 | -104.8181
19.7372 | -104.8205
19.7409 | -104.822
19.7446 | -104.8241
19.7483 | -104.8267
19.7519 | -104.8296
19.7556 | -104.8321
19.7593 | -104.8363
19.763 | -104.8387
19.7667 | -104.8381
19.7704 | -104.8373
19.7741 | -104.8361
19.7778 | -104.8363
19.7815 | -104.838
19.7851 | -104.8389
19.7888 | -104.8404
19.7925 | -104.8434
19.7962 | -104.8464
19.7999 | -104.8468
19.8036 | -104.8437
19.8073 | -104.8411
19.811 | -104.8417
19.8147 | -104.843
19.8183 | -104.8433
19.822 | -104.8419
19.8257 | -104.8405
19.8294 | -104.8388
19.8331 | -104.8374
19.8368 | -104.8346
19.8405 | -104.8291
19.8442 | -104.8233
19.8478 | -104.8195
19.8515 | -104.8158
19.8552 | -104.8122
19.8589 | -104.8085
19.8626 | -104.8053
19.8663 | -104.8025
19.87 | -104.8001
19.8737 | -104.7973
19.8774 | -104.7953
19.881 | -104.7938
19.8847 | -104.7914
19.8884 | -104.788
19.8921 | -104.7872
19.8958 | -104.7871
19.8995 | -104.7838
19.9032 | -104.7793
19.9069 | -104.7757
19.9106 | -104.7725
19.9142 | -104.7722
19.9179 | -104.771
19.9216 | -104.7687
19.9253 | -104.7694
19.929 | -104.7718
19.9327 | -104.7745
19.9364 | -104.778
19.9401 | -104.7807
19.9438 | -104.7817
19.9474 | -104.785
19.9511 | -104.7878
19.9548 | -104.7913
19.9585 | -104.7976
19.9622 | -104.8039
19.9659 | -104.8111
19.9696 | -104.8167
19.9733 | -104.8219
19.9769 | -104.8285
19.9806 | -104.8356
19.9843 | -104.8393
19.988 | -104.845
19.9917 | -104.8522
19.9954 | -104.8567
19.9991 | -104.861
20.0028 | -104.8634
20.0065 | -104.866
20.0101 | -104.8701
20.0138 | -104.8737
20.0175 | -104.8787
20.0212 | -104.882
20.0249 | -104.8848
20.0286 | -104.8852
20.0323 | -104.8872
20.036 | -104.8895
20.0397 | -104.8896
20.0433 | -104.8902
20.047 | -104.8912
20.0507 | -104.8902
20.0544 | -104.8914
20.0581 | -104.8936
20.0618 | -104.8935
20.0655 | -104.8914
20.0692 | -104.8881
20.0728 | -104.8823
20.0765 | -104.8784
20.0802 | -104.8742
20.0839 | -104.8717
20.0876 | -104.8693
20.0913 | -104.865
20.095 | -104.8606
20.0987 | -104.8533
20.1024 | -104.8477
20.106 | -104.8436
20.1097 | -104.8404
20.1134 | -104.8356
20.1171 | -104.8304
20.1208 | -104.8242
20.1245 | -104.8205
20.1282 | -104.8156
20.1319 | -104.8073
20.1356 | -104.7984
20.1392 | -104.7911
20.1429 | -104.7846
20.1466 | -104.7749
20.1503 | -104.7643
20.154 | -104.7527
20.1577 | -104.7434
20.1614 | -104.7343
20.1651 | -104.7233
20.1687 | -104.7128
20.1724 | -104.7051
20.1761 | -104.6957
20.1798 | -104.6862
20.1835 | -104.6778
20.1872 | -104.6696
20.1909 | -104.6597
20.1946 | -104.6503
20.1983 | -104.6422
20.2019 | -104.6327
20.2056 | -104.6204
20.2093 | -104.6075
20.213 | -104.5955
20.2167 | -104.585
20.2204 | -104.5708
20.2241 | -104.5596
20.2278 | -104.5496
20.2315 | -104.541
20.2351 | -104.5306
20.2388 | -104.521
20.2425 | -104.511
20.2462 | -104.4993
20.2499 | -104.4876
20.2536 | -104.475
20.2573 | -104.4662
20.261 | -104.4585
20.2646 | -104.4484
20.2683 | -104.4387
20.272 | -104.431
20.2757 | -104.4215
20.2794 | -104.4119
20.2831 | -104.4035
20.2868 | -104.3948
20.2905 | -104.3884
20.2942 | -104.3798
20.2978 | -104.3718
20.3015 | -104.3654
20.3052 | -104.3591
20.3089 | -104.3548
20.3126 | -104.3516
20.3163 | -104.3497
20.32 | -104.3467
20.3237 | -104.3447
20.3274 | -104.3434
20.331 | -104.3401
20.3347 | -104.3398
20.3384 | -104.3415
20.3421 | -104.3404
20.3458 | -104.34
20.3495 | -104.3391
20.3532 | -104.3373
20.3569 | -104.3388
20.3606 | -104.3408
20.3642 | -104.3392
20.3679 | -104.3393
20.3716 | -104.3391
20.3753 | -104.3373
20.379 | -104.336
20.3827 | -104.3333
20.3864 | -104.3309
20.3901 | -104.3274
20.3937 | -104.3229
20.3974 | -104.3189
20.4011 | -104.3142
20.4048 | -104.3068
20.4085 | -104.3015
20.4122 | -104.2964
20.4159 | -104.2933
20.4196 | -104.2889
20.4233 | -104.2813
20.4269 | -104.2732
20.4306 | -104.2662
20.4343 | -104.2596
20.438 | -104.2522
20.4417 | -104.2448
20.4454 | -104.2373
20.4491 | -104.23
20.4528 | -104.2236
20.4565 | -104.217
20.4601 | -104.2084
20.4638 | -104.1991
20.4675 | -104.1902
20.4712 | -104.181
20.4749 | -104.1719
20.4786 | -104.1631
20.4823 | -104.1534
20.486 | -104.142
20.4896 | -104.1319
20.4933 | -104.1237
20.497 | -104.1156
20.5007 | -104.1081
20.5044 | -104.1009
20.5081 | -104.0959
20.5118 | -104.0901
20.5155 | -104.0841
20.5192 | -104.076
20.5228 | -104.0685
20.5265 | -104.0657
20.5302 | -104.0622
20.5339 | -104.0602
20.5376 | -104.0603
20.5413 | -104.0615
20.545 | -104.0609
20.5487 | -104.0586
20.5524 | -104.058
20.556 | -104.0593
20.5597 | -104.0607
20.5634 | -104.0621
20.5671 | -104.0649
20.5708 | -104.0702
20.5745 | -104.0755
20.5782 | -104.0792
20.5819 | -104.0824
20.5856 | -104.0872
20.5892 | -104.0941
20.5929 | -104.099
20.5966 | -104.1027
20.6003 | -104.1071
20.604 | -104.1109
20.6077 | -104.113
20.6114 | -104.1146
20.6151 | -104.116
20.6187 | -104.1175
20.6224 | -104.1206
20.6261 | -104.121
20.6298 | -104.1206
20.6335 | -104.1227
20.6372 | -104.1232
20.6409 | -104.1258
20.6446 | -104.128
20.6483 | -104.1262
20.6519 | -104.1259
20.6556 | -104.127
20.6593 | -104.1276
20.663 | -104.1267
20.6667 | -104.1288
20.6704 | -104.1323
20.6741 | -104.1351
20.6778 | -104.1378
20.6815 | -104.139
20.6851 | -104.1395
20.6888 | -104.1404
20.6925 | -104.1442
20.6962 | -104.152
20.6999 | -104.1576
20.7036 | -104.1612
20.7073 | -104.1651
20.711 | -104.1716
20.7146 | -104.1803
20.7183 | -104.1878
20.722 | -104.1936
20.7257 | -104.2024
20.7294 | -104.2114
20.7331 | -104.2193
20.7368 | -104.2262
20.7405 | -104.2337
20.7442 | -104.2419
20.7478 | -104.2487
20.7515 | -104.2569
20.7552 | -104.2643
20.7589 | -104.2738
20.7626 | -104.2825
20.7663 | -104.2896
20.77 | -104.2977
20.7737 | -104.305
20.7774 | -104.312
20.781 | -104.3179
20.7847 | -104.3248
20.7884 | -104.3286
20.7921 | -104.332
20.7958 | -104.3343
20.7995 | -104.337
20.8032 | -104.3391
20.8069 | -104.3413
20.8105 | -104.343
20.8142 | -104.347
20.8179 | -104.3517
20.8216 | -104.3525
20.8253 | -104.3557
20.829 | -104.3587
20.8327 | -104.3592
20.8364 | -104.3604
20.8401 | -104.3642
20.8437 | -104.3675
20.8474 | -104.3694
20.8511 | -104.3733
20.8548 | -104.3769
20.8585 | -104.3809
20.8622 | -104.3825
20.8659 | -104.3824
20.8696 | -104.3838
20.8733 | -104.3889
20.8769 | -104.3935
20.8806 | -104.3951
20.8843 | -104.3964
20.888 | -104.3999
20.8917 | -104.4043
20.8954 | -104.4071
20.8991 | -104.4094
20.9028 | -104.4131
20.9064 | -104.4174
20.9101 | -104.421
20.9138 | -104.4254
20.9175 | -104.4294
20.9212 | -104.4333
20.9249 | -104.4373
20.9286 | -104.4413
20.9323 | -104.4451
20.936 | -104.4467
20.9396 | -104.4508
20.9433 | -104.4576
20.947 | -104.4621
20.9507 | -104.4663
20.9544 | -104.4698
20.9581 | -104.4742
20.9618 | -104.4789
20.9655 | -104.4834
20.9692 | -104.4884
20.9728 | -104.4951
20.9765 | -104.5032
20.9802 | -104.5124
20.9839 | -104.5182
20.9876 | -104.5212
20.9913 | -104.5232
20.995 | -104.5307
20.9987 | -104.5392
21.0024 | -104.5465
21.006 | -104.5533
21.0097 | -104.5609
21.0134 | -104.5691
21.0171 | -104.575
21.0208 | -104.5793
21.0245 | -104.5856
21.0282 | -104.5924
21.0319 | -104.6001
21.0355 | -104.6061
21.0392 | -104.6116
21.0429 | -104.6169
21.0466 | -104.6202
21.0503 | -104.6226
21.054 | -104.6247
21.0577 | -104.6247
21.0614 | -104.626
21.0651 | -104.6287
21.0687 | -104.6308
21.0724 | -104.6337
21.0761 | -104.6374
21.0798 | -104.6391
21.0835 | -104.6371
21.0872 | -104.6387
21.0909 | -104.6386
21.0946 | -104.6384
21.0983 | -104.6367
21.1019 | -104.6356
21.1056 | -104.6332
21.1093 | -104.6315
21.113 | -104.6304
21.1167 | -104.627
21.1204 | -104.6245
21.1241 | -104.6211
21.1278 | -104.6192
21.1314 | -104.6184
21.1351 | -104.6155
21.1388 | -104.6131
21.1425 | -104.6112
21.1462 | -104.6072
21.1499 | -104.6047
21.1536 | -104.6033
21.1573 | -104.6005
21.161 | -104.6014
21.1646 | -104.6024
21.1683 | -104.6021
21.172 | -104.5985
21.1757 | -104.5971
21.1794 | -104.5987
21.1831 | -104.5986
21.1868 | -104.5994
21.1905 | -104.6009
21.1942 | -104.6039
21.1978 | -104.6075
21.2015 | -104.6133
21.2052 | -104.6175
21.2089 | -104.6217
21.2126 | -104.6263
21.2163 | -104.6323
21.22 | -104.6399
21.2237 | -104.649
21.2274 | -104.6572
21.231 | -104.6666
21.2347 | -104.677
21.2384 | -104.6863
21.2421 | -104.6957
21.2458 | -104.7054
21.2495 | -104.7169
21.2532 | -104.7284
21.2569 | -104.7394
21.2605 | -104.7495
21.2642 | -104.7608
21.2679 | -104.7713
21.2716 | -104.7809
21.2753 | -104.7928
21.279 | -104.8036
21.2827 | -104.8136
21.2864 | -104.8265
21.2901 | -104.8403
21.2937 | -104.8536
21.2974 | -104.8635
21.3011 | -104.8708
21.3048 | -104.879
21.3085 | -104.8885
21.3122 | -104.898
21.3159 | -104.9065
21.3196 | -104.915
21.3233 | -104.925
21.3269 | -104.9335
21.3306 | -104.9433
21.3343 | -104.9512
21.338 | -104.9598
21.3417 | -104.9684
21.3454 | -104.9781
21.3491 | -104.9873
21.3528 | -104.9963
21.3564 | -105.0045
21.3601 | -105.0136
21.3638 | -105.0196
21.3675 | -105.0243
21.3712 | -105.0289
21.3749 | -105.0339
21.3786 | -105.0382
21.3823 | -105.0427
21.386 | -105.0469
21.3896 | -105.0506
21.3933 | -105.0544
21.397 | -105.0565
21.4007 | -105.0583
21.4044 | -105.0603
21.4081 | -105.0621
21.4118 | -105.0628
21.4155 | -105.0628
21.4192 | -105.0635
21.4228 | -105.0637
21.4265 | -105.063
21.4302 | -105.0603
21.4339 | -105.0595
21.4376 | -105.0581
21.4413 | -105.0562
21.445 | -105.0531
21.4487 | -105.0514
21.4523 | -105.0484
21.456 | -105.044
21.4597 | -105.038
21.4634 | -105.0308
21.4671 | -105.0254
21.4708 | -105.0196
21.4745 | -105.0137
21.4782 | -105.0067
21.4819 | -104.9989
21.4855 | -104.9918
21.4892 | -104.9845
21.4929 | -104.9768
21.4966 | -104.9694
21.5003 | -104.9606
21.504 | -104.9503
21.5077 | -104.9448
21.5114 | -104.9383
21.5151 | -104.9307
21.5187 | -104.9222
21.5224 | -104.9138
21.5261 | -104.9052
21.5298 | -104.8976
21.5335 | -104.8908
21.5372 | -104.8846
21.5409 | -104.8793
21.5446 | -104.8731
21.5482 | -104.8661
21.5519 | -104.859
21.5556 | -104.8532
21.5593 | -104.8479
21.563 | -104.8446
21.5667 | -104.8396
21.5704 | -104.8343
21.5741 | -104.8304
21.5778 | -104.8256
21.5814 | -104.8208
21.5851 | -104.8157
21.5888 | -104.809
21.5925 | -104.804
21.5962 | -104.7994
21.5999 | -104.7929
21.6036 | -104.7862
21.6073 | -104.7839
21.611 | -104.7797
21.6146 | -104.7726
21.6183 | -104.7652
21.622 | -104.7593
21.6257 | -104.7552
21.6294 | -104.748
21.6331 | -104.7412
21.6368 | -104.733
21.6405 | -104.7254
21.6442 | -104.7165
21.6478 | -104.7086
21.6515 | -104.6995
21.6552 | -104.6906
21.6589 | -104.6822
21.6626 | -104.6731
21.6663 | -104.6643
21.67 | -104.6554
21.6737 | -104.6455
21.6773 | -104.6341
21.681 | -104.6211
21.6847 | -104.6101
21.6884 | -104.5993
21.6921 | -104.5893
21.6958 | -104.5777
21.6995 | -104.5657
21.7032 | -104.5553
21.7069 | -104.5452
21.7105 | -104.5325
21.7142 | -104.519
21.7179 | -104.5046
21.7216 | -104.4882
21.7253 | -104.4736
21.729 | -104.4594
21.7327 | -104.4443
21.7364 | -104.4298
21.7401 | -104.4136
21.7437 | -104.3986
21.7474 | -104.3855
21.7511 | -104.3719
21.7548 | -104.3576
21.7585 | -104.3431
21.7622 | -104.3295
21.7659 | -104.3187
21.7696 | -104.3083
21.7732 | -104.2951
21.7769 | -104.2834
21.7806 | -104.2727
21.7843 | -104.2618
21.788 | -104.2498
21.7917 | -104.2372
21.7954 | -104.2262
21.7991 | -104.2158
21.8028 | -104.2054
21.8064 | -104.1973
21.8101 | -104.1883
21.8138 | -104.177
21.8175 | -104.1712
21.8212 | -104.1636
21.8249 | -104.1571
21.8286 | -104.1494
21.8323 | -104.1422
21.836 | -104.1349
21.8396 | -104.128
21.8433 | -104.1226
21.847 | -104.1173
21.8507 | -104.1122
21.8544 | -104.1063
21.8581 | -104.102
21.8618 | -104.0986
21.8655 | -104.0924
21.8692 | -104.0873
21.8728 | -104.0829
21.8765 | -104.0797
21.8802 | -104.0765
21.8839 | -104.0691
21.8876 | -104.0637
21.8913 | -104.0612
21.895 | -104.0573
21.8987 | -104.0526
21.9023 | -104.045
21.906 | -104.0394
21.9097 | -104.0361
21.9134 | -104.0318
21.9171 | -104.0259
21.9208 | -104.0204
21.9245 | -104.0172
21.9282 | -104.013
21.9319 | -104.0074
21.9355 | -104.0022
21.9392 | -103.9965
21.9429 | -103.9904
21.9466 | -103.9851
21.9503 | -103.9784
21.954 | -103.9711
21.9577 | -103.9639
21.9614 | -103.9567
21.9651 | -103.9498
21.9687 | -103.9406
21.9724 | -103.9326
21.9761 | -103.9225
21.9798 | -103.9139
21.9835 | -103.9071
21.9872 | -103.9
21.9909 | -103.8929
21.9946 | -103.8831
21.9982 | -103.8752
22.0019 | -103.8674
22.0056 | -103.8582
22.0093 | -103.8496
22.013 | -103.8417
22.0167 | -103.8324
22.0204 | -103.824
22.0241 | -103.8181
22.0278 | -103.8119
22.0314 | -103.8077
22.0351 | -103.8009
22.0388 | -103.7952
22.0425 | -103.793
22.0462 | -103.7905
22.0499 | -103.7875
22.0536 | -103.7861
22.0573 | -103.7829
22.061 | -103.7812
22.0646 | -103.7805
22.0683 | -103.7796
22.072 | -103.7789
22.0757 | -103.7774
22.0794 | -103.7761
22.0831 | -103.7792
22.0868 | -103.7814
22.0905 | -103.784
22.0941 | -103.7874
22.0978 | -103.7911
22.1015 | -103.7945
22.1052 | -103.7971
22.1089 | -103.8024
22.1126 | -103.8065
22.1163 | -103.8101
22.12 | -103.8146
22.1237 | -103.8185
22.1273 | -103.8224
22.131 | -103.8256
22.1347 | -103.8288
22.1384 | -103.83
22.1421 | -103.8309
22.1458 | -103.8309
22.1495 | -103.831
22.1532 | -103.8323
22.1569 | -103.8319
22.1605 | -103.8296
22.1642 | -103.8273
22.1679 | -103.8248
22.1716 | -103.8202
22.1753 | -103.8146
22.179 | -103.8117
22.1827 | -103.8048
22.1864 | -103.7969
22.19 | -103.7887
22.1937 | -103.7819
22.1974 | -103.777
22.2011 | -103.7696
22.2048 | -103.7603
22.2085 | -103.7532
22.2122 | -103.7466
22.2159 | -103.7406
22.2196 | -103.7299
22.2232 | -103.7197
22.2269 | -103.7121
22.2306 | -103.7032
22.2343 | -103.6942
22.238 | -103.6838
22.2417 | -103.6751
22.2454 | -103.6661
22.2491 | -103.6555
22.2528 | -103.6466
22.2564 | -103.6369
22.2601 | -103.6265
22.2638 | -103.6178
22.2675 | -103.6114
22.2712 | -103.6011
22.2749 | -103.5905
22.2786 | -103.5812
22.2823 | -103.5729
22.286 | -103.564
22.2896 | -103.5542
22.2933 | -103.5437
22.297 | -103.5339
22.3007 | -103.5253
22.3044 | -103.5162
22.3081 | -103.5062
22.3118 | -103.4952
22.3155 | -103.4853
22.3191 | -103.4766
22.3228 | -103.4692
22.3265 | -103.4613
22.3302 | -103.4557
22.3339 | -103.4495
22.3376 | -103.4427
22.3413 | -103.4381
22.345 | -103.4324
22.3487 | -103.429
22.3523 | -103.4239
22.356 | -103.4194
22.3597 | -103.4143
22.3634 | -103.4085
22.3671 | -103.4055
22.3708 | -103.4032
22.3745 | -103.4016
22.3782 | -103.4001
22.3819 | -103.3993
22.3855 | -103.3998
22.3892 | -103.4
22.3929 | -103.3983
22.3966 | -103.3976
22.4003 | -103.3965
22.404 | -103.3972
22.4077 | -103.3984
22.4114 | -103.3983
22.415 | -103.4
22.4187 | -103.4018
22.4224 | -103.4028
22.4261 | -103.4
22.4298 | -103.4003
22.4335 | -103.4008
22.4372 | -103.3999
22.4409 | -103.3995
22.4446 | -103.3965
22.4482 | -103.3916
22.4519 | -103.3892
22.4556 | -103.3849
22.4593 | -103.3798
22.463 | -103.3729
22.4667 | -103.365
22.4704 | -103.3501
22.4741 | -103.338
22.4778 | -103.3239
22.4814 | -103.3033
22.4851 | -103.2776
22.4888 | -103.2461
22.4925 | -103.2097
22.4962 | -103.1677
22.4999 | -103.1351
22.5036 | -103.1025
22.5073 | -103.0694
22.511 | -103.0307
22.5146 | -102.9906
22.5183 | -102.9498
22.522 | -102.9057
22.5257 | -102.8598
22.5294 | -102.8103
22.5331 | -102.7605
22.5368 | -102.7115
22.5405 | -102.6585
22.5441 | -102.6033
22.5478 | -102.5465
22.5515 | -102.4901
22.5552 | -102.4327
22.5589 | -102.3733
22.5626 | -102.3142
22.5663 | -102.257
22.57 | -102.1994
22.5737 | -102.1421
22.5773 | -102.085
22.581 | -102.0295
22.5847 | -101.9756
22.5884 | -101.9244
22.5921 | -101.8752
22.5958 | -101.8266
22.5995 | -101.7805
22.6032 | -101.7359
22.6069 | -101.6964
22.6105 | -101.6596
22.6142 | -101.6235
22.6179 | -101.5917
22.6216 | -101.5651
22.6253 | -101.5431
22.629 | -101.5219
22.6327 | -101.5042
22.6364 | -101.4919
22.64 | -101.4826
22.6437 | -101.4768
22.6474 | -101.4759
22.6511 | -101.4782
22.6548 | -101.4837
22.6585 | -101.4933
22.6622 | -101.5076
22.6659 | -101.5259
22.6696 | -101.5481
22.6732 | -101.5736
22.6769 | -101.6006
22.6806 | -101.6306
22.6843 | -101.6658
22.688 | -101.7021
22.6917 | -101.7407
22.6954 | -101.7821
22.6991 | -101.8249
22.7028 | -101.8688
22.7064 | -101.9139
22.7101 | -101.9615
22.7138 | -102.0118
22.7175 | -102.0615
22.7212 | -102.1116
22.7249 | -102.1616
22.7286 | -102.2117
22.7323 | -102.2616
22.7359 | -102.3091
22.7396 | -102.3563
22.7433 | -102.4027
22.747 | -102.4475
22.7507 | -102.4914
22.7544 | -102.5338
22.7581 | -102.5744
22.7618 | -102.6144
22.7655 | -102.6523
22.7691 | -102.6892
22.7728 | -102.7243
22.7765 | -102.7581
22.7802 | -102.7908
22.7839 | -102.8177
22.7876 | -102.8408
22.7913 | -102.8669
22.795 | -102.8876
22.7987 | -102.9092
22.8023 | -102.9284
22.806 | -102.9474
22.8097 | -102.968
22.8134 | -102.9851
22.8171 | -103.0015
22.8208 | -103.0048
22.8245 | -103.002
22.8282 | -102.9971
22.8318 | -102.9923
22.8355 | -102.9829
22.8392 | -102.9688
22.8429 | -102.9559
22.8466 | -102.9443
22.8503 | -102.9327
22.854 | -102.9154
22.8577 | -102.8983
22.8614 | -102.8835
22.865 | -102.8654
22.8687 | -102.8495
22.8724 | -102.8328
22.8761 | -102.815
22.8798 | -102.7954
22.8835 | -102.7778
22.8872 | -102.7597
22.8909 | -102.7405
22.8946 | -102.7186
22.8982 | -102.6983
22.9019 | -102.6781
22.9056 | -102.6554
22.9093 | -102.6338
22.913 | -102.6125
22.9167 | -102.5907
22.9204 | -102.5693
22.9241 | -102.5495
22.9278 | -102.5322
22.9314 | -102.5145
22.9351 | -102.498
22.9388 | -102.4826
22.9425 | -102.4664
22.9462 | -102.4502
22.9499 | -102.4341
22.9536 | -102.4211
22.9573 | -102.4079
22.9609 | -102.3934
22.9646 | -102.3794
22.9683 | -102.3665
22.972 | -102.3525
22.9757 | -102.3414
22.9794 | -102.3284
22.9831 | -102.314
22.9868 | -102.3005
22.9905 | -102.2869
22.9941 | -102.2729
22.9978 | -102.256
23.0015 | -102.2422
23.0052 | -102.2299
23.0089 | -102.2163
23.0126 | -102.2051
23.0163 | -102.1951
23.02 | -102.1841
23.0237 | -102.1737
23.0273 | -102.1632
23.031 | -102.1519
23.0347 | -102.1399
23.0384 | -102.1265
23.0421 | -102.116
23.0458 | -102.1064
23.0495 | -102.096
23.0532 | -102.0852
23.0568 | -102.0753
23.0605 | -102.0652
23.0642 | -102.055
23.0679 | -102.0456
23.0716 | -102.0371
23.0753 | -102.0263
23.079 | -102.0165
23.0827 | -102.0072
23.0864 | -101.9955
23.09 | -101.9849
23.0937 | -101.9749
23.0974 | -101.9628
23.1011 | -101.9538
23.1048 | -101.9445
23.1085 | -101.9386
23.1122 | -101.9333
23.1159 | -101.9257
23.1196 | -101.9172
23.1232 | -101.9071
23.1269 | -101.9005
23.1306 | -101.8929
23.1343 | -101.8856
23.138 | -101.8803
23.1417 | -101.8748
23.1454 | -101.8704
23.1491 | -101.8663
23.1528 | -101.8625
23.1564 | -101.861
23.1601 | -101.8612
23.1638 | -101.8633
23.1675 | -101.8662
23.1712 | -101.868
23.1749 | -101.8703
23.1786 | -101.8748
23.1823 | -101.8812
23.1859 | -101.8873
23.1896 | -101.8952
23.1933 | -101.9029
23.197 | -101.9119
23.2007 | -101.9227
23.2044 | -101.9345
23.2081 | -101.945
23.2118 | -101.9576
23.2155 | -101.9685
23.2191 | -101.9765
23.2228 | -101.9847
23.2265 | -101.9925
23.2302 | -102.001
23.2339 | -102.0125
23.2376 | -102.0219
23.2413 | -102.0334
23.245 | -102.0458
23.2487 | -102.0566
23.2523 | -102.0659
23.256 | -102.0751
23.2597 | -102.0829
23.2634 | -102.0885
23.2671 | -102.0945
23.2708 | -102.0997
23.2745 | -102.1044
23.2782 | -102.1102
23.2818 | -102.1163
23.2855 | -102.1231
23.2892 | -102.1282
23.2929 | -102.132
23.2966 | -102.1338
23.3003 | -102.1345
23.304 | -102.1334
23.3077 | -102.1349
23.3114 | -102.1369
23.315 | -102.1382
23.3187 | -102.1406
23.3224 | -102.1426
23.3261 | -102.1459
23.3298 | -102.1439
23.3335 | -102.1415
23.3372 | -102.1401
23.3409 | -102.1355
23.3446 | -102.1305
23.3482 | -102.1267
23.3519 | -102.1246
23.3556 | -102.1238
23.3593 | -102.1239
23.363 | -102.1251
23.3667 | -102.1285
23.3704 | -102.1287
23.3741 | -102.1281
23.3777 | -102.129
23.3814 | -102.131
23.3851 | -102.1348
23.3888 | -102.1398
23.3925 | -102.1447
23.3962 | -102.1485
23.3999 | -102.1517
23.4036 | -102.158
23.4073 | -102.1629
23.4109 | -102.1722
23.4146 | -102.1794
23.4183 | -102.1884
23.422 | -102.1959
23.4257 | -102.2038
23.4294 | -102.2119
23.4331 | -102.2184
23.4368 | -102.2254
23.4405 | -102.2307
23.4441 | -102.233
23.4478 | -102.2354
23.4515 | -102.2374
23.4552 | -102.243
23.4589 | -102.2484
23.4626 | -102.2551
23.4663 | -102.2594
23.47 | -102.2616
23.4736 | -102.2616
23.4773 | -102.2616
23.481 | -102.2635
23.4847 | -102.2656
23.4884 | -102.2662
23.4921 | -102.2662
23.4958 | -102.2636
23.4995 | -102.2584
23.5032 | -102.2483
23.5068 | -102.2391
23.5105 | -102.2304
23.5142 | -102.2182
23.5179 | -102.2043
23.5216 | -102.1919
23.5253 | -102.1799
23.529 | -102.1674
23.5327 | -102.1529
23.5364 | -102.1372
23.54 | -102.123
23.5437 | -102.1073
23.5474 | -102.0918
23.5511 | -102.079
23.5548 | -102.0642
23.5585 | -102.0485
23.5622 | -102.0322
23.5659 | -102.0118
23.5696 | -101.9908
23.5732 | -101.9678
23.5769 | -101.9452
23.5806 | -101.9256
23.5843 | -101.9062
23.588 | -101.8857
23.5917 | -101.8664
23.5954 | -101.8488
23.5991 | -101.8305
23.6027 | -101.814
23.6064 | -101.7985
23.6101 | -101.7836
23.6138 | -101.7706
23.6175 | -101.7574
23.6212 | -101.7452
23.6249 | -101.7364
23.6286 | -101.7294
23.6323 | -101.7243
23.6359 | -101.7215
23.6396 | -101.7185
23.6433 | -101.7157
23.647 | -101.7131
23.6507 | -101.7125
23.6544 | -101.7143
23.6581 | -101.7181
23.6618 | -101.7255
23.6655 | -101.7318
23.6691 | -101.7387
23.6728 | -101.749
23.6765 | -101.7594
23.6802 | -101.7718
23.6839 | -101.7842
23.6876 | -101.7978
23.6913 | -101.8132
23.695 | -101.8284
23.6986 | -101.8442
23.7023 | -101.8621
23.706 | -101.8806
23.7097 | -101.9008
23.7134 | -101.9217
23.7171 | -101.9421
23.7208 | -101.9627
23.7245 | -101.9838
23.7282 | -102.0064
23.7318 | -102.0302
23.7355 | -102.0522
23.7392 | -102.0757
23.7429 | -102.0998
23.7466 | -102.1231
23.7503 | -102.145
23.754 | -102.1678
23.7577 | -102.1928
23.7614 | -102.2163
23.765 | -102.2396
23.7687 | -102.2628
23.7724 | -102.2868
23.7761 | -102.3132
23.7798 | -102.3421
23.7835 | -102.3708
23.7872 | -102.3989
23.7909 | -102.423
23.7946 | -102.4449
23.7982 | -102.4666
23.8019 | -102.4893
23.8056 | -102.5119
23.8093 | -102.5358
23.813 | -102.5596
23.8167 | -102.5819
23.8204 | -102.6028
23.8241 | -102.6244
23.8277 | -102.6463
23.8314 | -102.6697
23.8351 | -102.6939
23.8388 | -102.721
23.8425 | -102.7475
23.8462 | -102.7716
23.8499 | -102.7947
23.8536 | -102.8183
23.8573 | -102.8414
23.8609 | -102.8621
23.8646 | -102.8816
23.8683 | -102.9023
23.872 | -102.9239
23.8757 | -102.9456
23.8794 | -102.9679
23.8831 | -102.989
23.8868 | -103.0088
23.8905 | -103.0324
23.8941 | -103.0539
23.8978 | -103.0731
23.9015 | -103.0942
23.9052 | -103.1143
23.9089 | -103.1329
23.9126 | -103.1508
23.9163 | -103.1662
23.92 | -103.1814
23.9236 | -103.1983
23.9273 | -103.2153
23.931 | -103.2331
23.9347 | -103.2514
23.9384 | -103.2676
23.9421 | -103.2841
23.9458 | -103.2951
23.9495 | -103.3093
23.9532 | -103.3225
23.9568 | -103.336
23.9605 | -103.3512
23.9642 | -103.3652
23.9679 | -103.3803
23.9716 | -103.3938
23.9753 | -103.4034
23.979 | -103.4136
23.9827 | -103.4259
23.9864 | -103.4367
23.99 | -103.4421
23.9937 | -103.4475
23.9974 | -103.455
24.0011 | -103.4626
24.0048 | -103.4714
24.0085 | -103.4785
24.0122 | -103.484
24.0159 | -103.4904
24.0195 | -103.4946
24.0232 | -103.4969
24.0269 | -103.4985
24.0306 | -103.4969
24.0343 | -103.4963
24.038 | -103.4965
24.0417 | -103.4961
24.0454 | -103.495
24.0491 | -103.4928
24.0527 | -103.4904
24.0564 | -103.4888
24.0601 | -103.4875
24.0638 | -103.4838
24.0675 | -103.4792
24.0712 | -103.4719
24.0749 | -103.4611
24.0786 | -103.4521
24.0823 | -103.4435
24.0859 | -103.4344
24.0896 | -103.4247
24.0933 | -103.4139
24.097 | -103.4021
24.1007 | -103.3912
24.1044 | -103.3783
24.1081 | -103.3643
24.1118 | -103.3504
24.1154 | -103.3366
24.1191 | -103.3235
24.1228 | -103.3079
24.1265 | -103.2922
24.1302 | -103.2775
24.1339 | -103.2639
24.1376 | -103.25
24.1413 | -103.2352
24.145 | -103.2206
24.1486 | -103.2064
24.1523 | -103.1918
24.156 | -103.1777
24.1597 | -103.1623
24.1634 | -103.1481
24.1671 | -103.1348
24.1708 | -103.1207
24.1745 | -103.1079
24.1782 | -103.095
24.1818 | -103.0827
24.1855 | -103.0706
24.1892 | -103.0563
24.1929 | -103.0423
24.1966 | -103.0293
24.2003 | -103.018
24.204 | -103.0077
24.2077 | -102.9978
24.2114 | -102.9866
24.215 | -102.9728
24.2187 | -102.9638
24.2224 | -102.9583
24.2261 | -102.9512
24.2298 | -102.9394
24.2335 | -102.926
24.2372 | -102.9129
24.2409 | -102.9
24.2445 | -102.8798
24.2482 | -102.8498
24.2519 | -102.8141
24.2556 | -102.7756
24.2593 | -102.7321
24.263 | -102.6807
24.2667 | -102.6323
24.2704 | -102.5897
24.2741 | -102.5499
24.2777 | -102.51
24.2814 | -102.4676
24.2851 | -102.4233
24.2888 | -102.3777
24.2925 | -102.3289
24.2962 | -102.2767
24.2999 | -102.2224
24.3036 | -102.1656
24.3073 | -102.1072
24.3109 | -102.0442
24.3146 | -101.9789
24.3183 | -101.9163
24.322 | -101.8506
24.3257 | -101.7829
24.3294 | -101.7145
24.3331 | -101.6466
24.3368 | -101.5772
24.3404 | -101.5067
24.3441 | -101.4359
24.3478 | -101.3651
24.3515 | -101.2949
24.3552 | -101.2257
24.3589 | -101.1566
24.3626 | -101.0901
24.3663 | -101.0263
24.37 | -100.9648
24.3736 | -100.907
24.3773 | -100.8511
24.381 | -100.7955
24.3847 | -100.7425
24.3884 | -100.6942
24.3921 | -100.6494
24.3958 | -100.6085
24.3995 | -100.5711
24.4032 | -100.5373
24.4068 | -100.5093
24.4105 | -100.4855
24.4142 | -100.4667
24.4179 | -100.4528
24.4216 | -100.4414
24.4253 | -100.434
24.429 | -100.4313
24.4327 | -100.4329
24.4364 | -100.4396
24.44 | -100.4509
24.4437 | -100.467
24.4474 | -100.4883
24.4511 | -100.5129
24.4548 | -100.5412
24.4585 | -100.5732
24.4622 | -100.6082
24.4659 | -100.6481
24.4695 | -100.6918
24.4732 | -100.7384
24.4769 | -100.8077
24.4806 | -100.8574
24.4843 | -100.9095
24.488 | -100.9648
24.4917 | -101.0208
24.4954 | -101.0784
24.4991 | -101.1384
24.5027 | -101.1998
24.5064 | -101.2633
24.5101 | -101.3261
24.5138 | -101.39
24.5175 | -101.4539
24.5212 | -101.5172
24.5249 | -101.581
24.5286 | -101.6445
24.5323 | -101.7093
24.5359 | -101.7747
24.5396 | -101.8388
24.5433 | -101.901
24.547 | -101.9614
24.5507 | -102.02
24.5544 | -102.0781
24.5581 | -102.1343
24.5618 | -102.1894
24.5654 | -102.2349
24.5691 | -102.2711
24.5728 | -102.3063
24.5765 | -102.3365
24.5802 | -102.3371
24.5839 | -102.3241
24.5876 | -102.2998
24.5913 | -102.2604
24.595 | -102.2465
24.5986 | -102.2396
24.6023 | -102.2323
24.606 | -102.2193
24.6097 | -102.2023
24.6134 | -102.1714
24.6171 | -102.1252
24.6208 | -102.0683
24.6245 | -102.0037
24.6282 | -101.9325
24.6318 | -101.8552
24.6355 | -101.7754
24.6392 | -101.6929
24.6429 | -101.6095
24.6466 | -101.5238
24.6503 | -101.4332
24.654 | -101.3403
24.6577 | -101.2469
24.6613 | -101.152
24.665 | -101.0538
24.6687 | -100.9558
24.6724 | -100.8561
24.6761 | -100.7561
24.6798 | -100.6566
24.6835 | -100.5573
24.6872 | -100.4586
24.6909 | -100.3605
24.6945 | -100.2643
24.6982 | -100.1722
24.7019 | -100.0828
24.7056 | -99.9965
24.7093 | -99.9137
24.713 | -99.8348
24.7167 | -99.7596
24.7204 | -99.6883
24.7241 | -99.621
24.7277 | -99.5603
24.7314 | -99.5047
24.7351 | -99.4533
24.7388 | -99.4072
24.7425 | -99.3671
24.7462 | -99.3322
24.7499 | -99.3033
24.7536 | -99.2797
24.7572 | -99.2623
24.7609 | -99.2512
24.7646 | -99.2462
24.7683 | -99.2464
24.772 | -99.253
24.7757 | -99.2659
24.7794 | -99.2853
24.7831 | -99.3103
24.7868 | -99.3409
24.7904 | -99.3763
24.7941 | -99.4175
24.7978 | -99.4651
24.8015 | -99.5182
24.8052 | -99.5752
24.8089 | -99.6367
24.8126 | -99.7033
24.8163 | -99.7746
24.82 | -99.8501
24.8236 | -99.9291
24.8273 | -100.0131
24.831 | -100.1014
24.8347 | -100.1924
24.8384 | -100.2867
24.8421 | -100.3848
24.8458 | -100.4844
24.8495 | -100.585
24.8532 | -100.6878
24.8568 | -100.7917
24.8605 | -100.8971
24.8642 | -101.0017
24.8679 | -101.1056
24.8716 | -101.2087
24.8753 | -101.3108
24.879 | -101.4113
24.8827 | -101.5098
24.8863 | -101.6073
24.89 | -101.7012
24.8937 | -101.7907
24.8974 | -101.878
24.9011 | -101.9622
24.9048 | -102.0412
24.9085 | -102.1181
24.9122 | -102.1927
24.9159 | -102.2674
24.9195 | -102.3435
24.9232 | -102.4135
24.9269 | -102.4789
24.9306 | -102.5647
24.9343 | -102.649
24.938 | -102.7269
24.9417 | -102.8089
24.9454 | -102.8498
24.9491 | -102.8727
24.9527 | -102.8851
24.9564 | -102.8934
24.9601 | -102.9
24.9638 | -102.9071
24.9675 | -102.9132
24.9712 | -102.9172
24.9749 | -102.9214
24.9786 | -102.925
24.9822 | -102.9296
24.9859 | -102.9298
24.9896 | -102.928
24.9933 | -102.9235
24.997 | -102.9157
25.0007 | -102.9102
25.0044 | -102.9048
25.0081 | -102.9037
25.0118 | -102.8991
25.0154 | -102.8934
25.0191 | -102.8857
25.0228 | -102.8787
25.0265 | -102.8748
25.0302 | -102.8708
25.0339 | -102.8678
25.0376 | -102.8646
25.0413 | -102.8618
25.045 | -102.861
25.0486 | -102.8606
25.0523 | -102.8593
25.056 | -102.8587
25.0597 | -102.8587
25.0634 | -102.8594
25.0671 | -102.8598
25.0708 | -102.861
25.0745 | -102.8625
25.0782 | -102.8676
25.0818 | -102.8721
25.0855 | -102.8758
25.0892 | -102.8795
25.0929 | -102.8848
25.0966 | -102.8904
25.1003 | -102.8987
25.104 | -102.9062
25.1077 | -102.9138
25.1113 | -102.9217
25.115 | -102.9299
25.1187 | -102.9391
25.1224 | -102.948
25.1261 | -102.9546
25.1298 | -102.9626
25.1335 | -102.9707
25.1372 | -102.9766
25.1409 | -102.9822
25.1445 | -102.9913
25.1482 | -102.9986
25.1519 | -103.0047
25.1556 | -103.009
25.1593 | -103.0144
25.163 | -103.0226
25.1667 | -103.0305
25.1704 | -103.035
25.1741 | -103.039
25.1777 | -103.0424
25.1814 | -103.0429
25.1851 | -103.0415
25.1888 | -103.042
25.1925 | -103.0404
25.1962 | -103.0347
25.1999 | -103.0312
25.2036 | -103.0255
25.2072 | -103.0197
25.2109 | -103.0141
25.2146 | -103.0072
25.2183 | -103.0005
25.222 | -102.9937
25.2257 | -102.9874
25.2294 | -102.9809
25.2331 | -102.9726
25.2368 | -102.9626
25.2404 | -102.9522
25.2441 | -102.9421
25.2478 | -102.9356
25.2515 | -102.9262
25.2552 | -102.9143
25.2589 | -102.9036
25.2626 | -102.8888
25.2663 | -102.8676
25.27 | -102.8487
25.2736 | -102.8302
25.2773 | -102.8118
25.281 | -102.7963
25.2847 | -102.7796
25.2884 | -102.762
25.2921 | -102.743
25.2958 | -102.7231
25.2995 | -102.7042
25.3031 | -102.6824
25.3068 | -102.6604
25.3105 | -102.638
25.3142 | -102.6163
25.3179 | -102.5946
25.3216 | -102.573
25.3253 | -102.5527
25.329 | -102.5314
25.3327 | -102.5116
25.3363 | -102.491
25.34 | -102.4711
25.3437 | -102.451
25.3474 | -102.4299
25.3511 | -102.41
25.3548 | -102.3875
25.3585 | -102.3636
25.3622 | -102.3418
25.3659 | -102.3149
25.3695 | -102.2886
25.3732 | -102.2649
25.3769 | -102.245
25.3806 | -102.2254
25.3843 | -102.2021
25.388 | -102.1778
25.3917 | -102.1558
25.3954 | -102.1356
25.399 | -102.1123
25.4027 | -102.0881
25.4064 | -102.0665
25.4101 | -102.044
25.4138 | -102.0197
25.4175 | -101.994
25.4212 | -101.9697
25.4249 | -101.9475
25.4286 | -101.9245
25.4322 | -101.9023
25.4359 | -101.8782
25.4396 | -101.8531
25.4433 | -101.8268
25.447 | -101.802
25.4507 | -101.776
25.4544 | -101.7508
25.4581 | -101.725
25.4618 | -101.6984
25.4654 | -101.6725
25.4691 | -101.649
25.4728 | -101.6236
25.4765 | -101.5978
25.4802 | -101.573
25.4839 | -101.5457
25.4876 | -101.5185
25.4913 | -101.491
25.495 | -101.4649
25.4986 | -101.4389
25.5023 | -101.4062
25.506 | -101.3757
25.5097 | -101.3493
25.5134 | -101.3241
25.5171 | -101.2987
25.5208 | -101.2725
25.5245 | -101.2455
25.5281 | -101.2194
25.5318 | -101.1937
25.5355 | -101.169
25.5392 | -101.1497
25.5429 | -101.1277
25.5466 | -101.1064
25.5503 | -101.0835
25.554 | -101.0628
25.5577 | -101.0458
25.5613 | -101.0286
25.565 | -101.0132
25.5687 | -100.9973
25.5724 | -100.9813
25.5761 | -100.9674
25.5798 | -100.955
25.5835 | -100.9436
25.5872 | -100.9337
25.5909 | -100.9244
25.5945 | -100.9166
25.5982 | -100.9092
25.6019 | -100.9036
25.6056 | -100.8999
25.6093 | -100.8955
25.613 | -100.8931
25.6167 | -100.8936
25.6204 | -100.8946
25.624 | -100.899
25.6277 | -100.9067
25.6314 | -100.9143
25.6351 | -100.9217
25.6388 | -100.9297
25.6425 | -100.9383
25.6462 | -100.948
25.6499 | -100.9584
25.6536 | -100.9722
25.6572 | -100.9878
25.6609 | -101.0029
25.6646 | -101.0193
25.6683 | -101.0361
25.672 | -101.0556
25.6757 | -101.0749
25.6794 | -101.094
25.6831 | -101.1145
25.6868 | -101.1357
25.6904 | -101.1573
25.6941 | -101.1783
25.6978 | -101.2012
25.7015 | -101.2242
25.7052 | -101.2482
25.7089 | -101.2729
25.7126 | -101.2978
25.7163 | -101.3247
25.72 | -101.3521
25.7236 | -101.3776
25.7273 | -101.4066
25.731 | -101.4357
25.7347 | -101.4621
25.7384 | -101.4865
25.7421 | -101.5106
25.7458 | -101.5363
25.7495 | -101.5629
25.7531 | -101.5876
25.7568 | -101.612
25.7605 | -101.6394
25.7642 | -101.6663
25.7679 | -101.6903
25.7716 | -101.7136
25.7753 | -101.7388
25.779 | -101.7649
25.7827 | -101.7884
25.7863 | -101.8108
25.79 | -101.8343
25.7937 | -101.8563
25.7974 | -101.8795
25.8011 | -101.9027
25.8048 | -101.927
25.8085 | -101.9505
25.8122 | -101.9733
25.8159 | -101.9934
25.8195 | -102.0134
25.8232 | -102.0357
25.8269 | -102.0567
25.8306 | -102.0745
25.8343 | -102.0938
25.838 | -102.113
25.8417 | -102.1329
25.8454 | -102.1553
25.849 | -102.1734
25.8527 | -102.1913
25.8564 | -102.2088
25.8601 | -102.2268
25.8638 | -102.2506
25.8675 | -102.2707
25.8712 | -102.2873
25.8749 | -102.3034
25.8786 | -102.3196
25.8822 | -102.3359
25.8859 | -102.3535
25.8896 | -102.3714
25.8933 | -102.3876
25.897 | -102.4047
25.9007 | -102.4216
25.9044 | -102.437
25.9081 | -102.4534
25.9118 | -102.4709
25.9154 | -102.49
25.9191 | -102.5059
25.9228 | -102.521
25.9265 | -102.5363
25.9302 | -102.5519
25.9339 | -102.5658
25.9376 | -102.5785
25.9413 | -102.5918
25.9449 | -102.6036
25.9486 | -102.6154
25.9523 | -102.6205
25.956 | -102.626
25.9597 | -102.6311
25.9634 | -102.6379
25.9671 | -102.6455
25.9708 | -102.6528
25.9745 | -102.6569
25.9781 | -102.6576
25.9818 | -102.6587
25.9855 | -102.6583
25.9892 | -102.657
25.9929 | -102.6554
25.9966 | -102.651
26.0003 | -102.6473
26.004 | -102.643
26.0077 | -102.6387
26.0113 | -102.6324
26.015 | -102.6257
26.0187 | -102.6152
26.0224 | -102.6017
26.0261 | -102.5911
26.0298 | -102.5792
26.0335 | -102.5672
26.0372 | -102.5526
26.0408 | -102.5377
26.0445 | -102.5227
26.0482 | -102.5079
26.0519 | -102.4929
26.0556 | -102.4775
26.0593 | -102.4628
26.063 | -102.4459
26.0667 | -102.4273
26.0704 | -102.4075
26.074 | -102.3876
26.0777 | -102.3677
26.0814 | -102.3485
26.0851 | -102.3309
26.0888 | -102.3135
26.0925 | -102.2966
26.0962 | -102.2811
26.0999 | -102.264
26.1036 | -102.2498
26.1072 | -102.235
26.1109 | -102.2201
26.1146 | -102.2056
26.1183 | -102.1922
26.122 | -102.1786
26.1257 | -102.1634
26.1294 | -102.1492
26.1331 | -102.1377
26.1368 | -102.1262
26.1404 | -102.1173
26.1441 | -102.1089
26.1478 | -102.1012
26.1515 | -102.0944
26.1552 | -102.0879
26.1589 | -102.083
26.1626 | -102.0806
26.1663 | -102.0776
26.1699 | -102.0748
26.1736 | -102.0737
26.1773 | -102.0702
26.181 | -102.067
26.1847 | -102.0674
26.1884 | -102.0687
26.1921 | -102.0705
26.1958 | -102.0727
26.1995 | -102.075
26.2031 | -102.0785
26.2068 | -102.0819
26.2105 | -102.0853
26.2142 | -102.0902
26.2179 | -102.0968
26.2216 | -102.1043
26.2253 | -102.1109
26.229 | -102.1169
26.2327 | -102.1237
26.2363 | -102.1305
26.24 | -102.1363
26.2437 | -102.1419
26.2474 | -102.1465
26.2511 | -102.1506
26.2548 | -102.1549
26.2585 | -102.1601
26.2622 | -102.1661
26.2658 | -102.1696
26.2695 | -102.1748
26.2732 | -102.1817
26.2769 | -102.1863
26.2806 | -102.1898
26.2843 | -102.1939
26.288 | -102.1972
26.2917 | -102.1999
26.2954 | -102.2021
26.299 | -102.2042
26.3027 | -102.2097
26.3064 | -102.2161
26.3101 | -102.218
26.3138 | -102.2182
26.3175 | -102.216
26.3212 | -102.2196
26.3249 | -102.2243
26.3286 | -102.2278
26.3322 | -102.2289
26.3359 | -102.2279
26.3396 | -102.2266
26.3433 | -102.2278
26.347 | -102.231
26.3507 | -102.2313
26.3544 | -102.2308
26.3581 | -102.2282
26.3618 | -102.2245
26.3654 | -102.2182
26.3691 | -102.2151
26.3728 | -102.2118
26.3765 | -102.207
26.3802 | -102.2037
26.3839 | -102.2001
26.3876 | -102.197
26.3913 | -102.1947
26.3949 | -102.1898
26.3986 | -102.184
26.4023 | -102.1751
26.406 | -102.1646
26.4097 | -102.1555
26.4134 | -102.1457
26.4171 | -102.1352
26.4208 | -102.1256
26.4245 | -102.1179
26.4281 | -102.1098
26.4318 | -102.1015
26.4355 | -102.0941
26.4392 | -102.0864
26.4429 | -102.079
26.4466 | -102.0726
26.4503 | -102.0661
26.454 | -102.0592
26.4577 | -102.0504
26.4613 | -102.0417
26.465 | -102.0319
26.4687 | -102.0209
26.4724 | -102.0116
26.4761 | -102.0023
26.4798 | -101.9934
26.4835 | -101.9846
26.4872 | -101.9758
26.4908 | -101.9663
26.4945 | -101.9568
26.4982 | -101.9496
26.5019 | -101.9394
26.5056 | -101.9292
26.5093 | -101.9213
26.513 | -101.9118
26.5167 | -101.9038
26.5204 | -101.8973
26.524 | -101.8907
26.5277 | -101.884
26.5314 | -101.8775
26.5351 | -101.8717
26.5388 | -101.8662
26.5425 | -101.8618
26.5462 | -101.8567
26.5499 | -101.8555
26.5536 | -101.8555
26.5572 | -101.8532
26.5609 | -101.8511
26.5646 | -101.8495
26.5683 | -101.8492
26.572 | -101.8495
26.5757 | -101.85
26.5794 | -101.8526
26.5831 | -101.8552
26.5867 | -101.8569
26.5904 | -101.8588
26.5941 | -101.8605
26.5978 | -101.865
26.6015 | -101.8711
26.6052 | -101.8847
26.6089 | -101.891
26.6126 | -101.8985
26.6163 | -101.9065
26.6199 | -101.9158
26.6236 | -101.9269
26.6273 | -101.9372
26.631 | -101.9463
26.6347 | -101.9552
26.6384 | -101.9624
26.6421 | -101.9697
26.6458 | -101.9775
26.6495 | -101.9836
26.6531 | -101.9909
26.6568 | -101.9995
26.6605 | -102.0075
26.6642 | -102.0179
26.6679 | -102.0288
26.6716 | -102.0386
26.6753 | -102.0448
26.679 | -102.0513
26.6826 | -102.0538
26.6863 | -102.0577
26.69 | -102.0599
26.6937 | -102.0657
26.6974 | -102.0695
26.7011 | -102.0675
26.7048 | -102.0651
26.7085 | -102.0634
26.7122 | -102.0598
26.7158 | -102.0557
26.7195 | -102.0505
26.7232 | -102.0437
26.7269 | -102.0373
26.7306 | -102.0311
26.7343 | -102.0228
26.738 | -102.0125
26.7417 | -102.0031
26.7454 | -101.9877
26.749 | -101.9729
26.7527 | -101.9569
26.7564 | -101.9376
26.7601 | -101.9161
26.7638 | -101.8929
26.7675 | -101.8689
26.7712 | -101.844
26.7749 | -101.8156
26.7786 | -101.7868
26.7822 | -101.7612
26.7859 | -101.7348
26.7896 | -101.7047
26.7933 | -101.6733
26.797 | -101.6391
26.8007 | -101.6006
26.8044 | -101.5616
26.8081 | -101.524
26.8117 | -101.4824
26.8154 | -101.4417
26.8191 | -101.4027
26.8228 | -101.3619
26.8265 | -101.3158
26.8302 | -101.2707
26.8339 | -101.2323
26.8376 | -101.1912
26.8413 | -101.1483
26.8449 | -101.106
26.8486 | -101.0648
26.8523 | -101.0228
26.856 | -100.9792
26.8597 | -100.9353
26.8634 | -100.8921
26.8671 | -100.8505
26.8708 | -100.8093
26.8745 | -100.7669
26.8781 | -100.7257
26.8818 | -100.6835
26.8855 | -100.6394
26.8892 | -100.5981
26.8929 | -100.5576
26.8966 | -100.5176
26.9003 | -100.4783
26.904 | -100.44
26.9076 | -100.4027
26.9113 | -100.3647
26.915 | -100.3274
26.9187 | -100.2925
26.9224 | -100.2575
26.9261 | -100.2242
26.9298 | -100.1918
26.9335 | -100.1621
26.9372 | -100.1333
26.9408 | -100.1061
26.9445 | -100.0806
26.9482 | -100.0558
26.9519 | -100.0326
26.9556 | -100.0103
26.9593 | -99.991
26.963 | -99.9731
26.9667 | -99.9569
26.9704 | -99.9428
26.974 | -99.9299
26.9777 | -99.9184
26.9814 | -99.9079
26.9851 | -99.8994
26.9888 | -99.8944
26.9925 | -99.8916
26.9962 | -99.8901
26.9999 | -99.8909
27.0036 | -99.8941
27.0072 | -99.9004
27.0109 | -99.9094
27.0146 | -99.9204
27.0183 | -99.9332
27.022 | -99.9489
27.0257 | -99.9655
27.0294 | -99.9847
27.0331 | -100.0046
27.0367 | -100.0243
27.0404 | -100.0457
27.0441 | -100.0684
27.0478 | -100.0924
27.0515 | -100.1192
27.0552 | -100.1491
27.0589 | -100.1793
27.0626 | -100.2125
27.0663 | -100.2461
27.0699 | -100.2832
27.0736 | -100.3192
27.0773 | -100.355
27.081 | -100.3915
27.0847 | -100.4239
27.0884 | -100.4549
27.0921 | -100.4854
27.0958 | -100.5121
27.0995 | -100.542
27.1031 | -100.5753
27.1068 | -100.6073
27.1105 | -100.6384
27.1142 | -100.6708
27.1179 | -100.7039
27.1216 | -100.7362
27.1253 | -100.7673
27.129 | -100.7964
27.1326 | -100.8254
27.1363 | -100.8367
27.14 | -100.8621
27.1437 | -100.8866
27.1474 | -100.9103
27.1511 | -100.9322
27.1548 | -100.9507
27.1585 | -100.9681
27.1622 | -100.984
27.1658 | -100.9942
27.1695 | -101.0017
27.1732 | -101.0102
27.1769 | -101.0151
27.1806 | -101.0185
27.1843 | -101.0215
27.188 | -101.0225
27.1917 | -101.0213
27.1954 | -101.0212
27.199 | -101.0165
27.2027 | -101.0082
27.2064 | -101.0005
27.2101 | -100.996
27.2138 | -100.9891
27.2175 | -100.975
27.2212 | -100.963
27.2249 | -100.9537
27.2285 | -100.9429
27.2322 | -100.9307
27.2359 | -100.9183
27.2396 | -100.9077
27.2433 | -100.8974
27.247 | -100.8861
27.2507 | -100.8733
27.2544 | -100.8608
27.2581 | -100.8508
27.2617 | -100.8437
27.2654 | -100.8376
27.2691 | -100.8324
27.2728 | -100.8257
27.2765 | -100.8191
27.2802 | -100.8145
27.2839 | -100.811
27.2876 | -100.8096
27.2913 | -100.8097
27.2949 | -100.8124
27.2986 | -100.817
27.3023 | -100.8211
27.306 | -100.8275
27.3097 | -100.8334
27.3134 | -100.84
27.3171 | -100.8471
27.3208 | -100.8556
27.3244 | -100.8641
27.3281 | -100.8738
27.3318 | -100.8857
27.3355 | -100.8986
27.3392 | -100.9128
27.3429 | -100.9274
27.3466 | -100.9423
27.3503 | -100.9564
27.354 | -100.9709
27.3576 | -100.9861
27.3613 | -101.0014
27.365 | -101.0194
27.3687 | -101.037
27.3724 | -101.0527
27.3761 | -101.0683
27.3798 | -101.0846
27.3835 | -101.0995
27.3872 | -101.1139
27.3908 | -101.1269
27.3945 | -101.1383
27.3982 | -101.1507
27.4019 | -101.1643
27.4056 | -101.1759
27.4093 | -101.1878
27.413 | -101.1989
27.4167 | -101.2095
27.4204 | -101.2218
27.424 | -101.2334
27.4277 | -101.2438
27.4314 | -101.2535
27.4351 | -101.2622
27.4388 | -101.2685
27.4425 | -101.2746
27.4462 | -101.2809
27.4499 | -101.2871
27.4535 | -101.2913
27.4572 | -101.2952
27.4609 | -101.2995
27.4646 | -101.3043
27.4683 | -101.3129
27.472 | -101.321
27.4757 | -101.3279
27.4794 | -101.3373
27.4831 | -101.342
27.4867 | -101.3413
27.4904 | -101.341
27.4941 | -101.3417
27.4978 | -101.3413
27.5015 | -101.3392
27.5052 | -101.3371
27.5089 | -101.3352
27.5126 | -101.3328
27.5163 | -101.3298
27.5199 | -101.3254
27.5236 | -101.3225
27.5273 | -101.3204
27.531 | -101.32
27.5347 | -101.3179
27.5384 | -101.3135
27.5421 | -101.3063
27.5458 | -101.2979
27.5494 | -101.2896
27.5531 | -101.2801
27.5568 | -101.27
27.5605 | -101.2642
27.5642 | -101.2614
27.5679 | -101.2585
27.5716 | -101.2547
27.5753 | -101.2512
27.579 | -101.2467
27.5826 | -101.2426
27.5863 | -101.2395
27.59 | -101.2357
27.5937 | -101.2307
27.5974 | -101.2254
27.6011 | -101.2192
27.6048 | -101.2129
27.6085 | -101.2056
27.6122 | -101.1975
27.6158 | -101.1898
27.6195 | -101.1804
27.6232 | -101.1709
27.6269 | -101.1611
27.6306 | -101.1518
27.6343 | -101.1424
27.638 | -101.1339
27.6417 | -101.1243
27.6454 | -101.1116
27.649 | -101.0984
27.6527 | -101.0845
27.6564 | -101.071
27.6601 | -101.0588
27.6638 | -101.0468
27.6675 | -101.0355
27.6712 | -101.0219
27.6749 | -101.0084
27.6785 | -100.9945
27.6822 | -100.9811
27.6859 | -100.9682
27.6896 | -100.9571
27.6933 | -100.9452
27.697 | -100.9313
27.7007 | -100.917
27.7044 | -100.906
27.7081 | -100.8958
27.7117 | -100.8877
27.7154 | -100.8808
27.7191 | -100.8738
27.7228 | -100.8676
27.7265 | -100.8622
27.7302 | -100.8567
27.7339 | -100.8522
27.7376 | -100.8503
27.7413 | -100.8494
27.7449 | -100.8486
27.7486 | -100.8493
27.7523 | -100.8508
27.756 | -100.8531
27.7597 | -100.8555
27.7634 | -100.8585
27.7671 | -100.8626
27.7708 | -100.8674
27.7744 | -100.8718
27.7781 | -100.8789
27.7818 | -100.8866
27.7855 | -100.8954
27.7892 | -100.9047
27.7929 | -100.914
27.7966 | -100.9247
27.8003 | -100.9346
27.804 | -100.9453
27.8076 | -100.9564
27.8113 | -100.9688
27.815 | -100.9809
27.8187 | -100.9923
27.8224 | -101.005
27.8261 | -101.0169
27.8298 | -101.031
27.8335 | -101.0448
27.8372 | -101.0573
27.8408 | -101.0694
27.8445 | -101.0816
27.8482 | -101.0921
27.8519 | -101.1048
27.8556 | -101.1179
27.8593 | -101.1311
27.863 | -101.145
27.8667 | -101.1568
27.8703 | -101.169
27.874 | -101.1816
27.8777 | -101.1934
27.8814 | -101.2053
27.8851 | -101.2162
27.8888 | -101.2272
27.8925 | -101.2402
27.8962 | -101.2528
27.8999 | -101.2656
27.9035 | -101.2783
27.9072 | -101.2907
27.9109 | -101.3034
27.9146 | -101.3152
27.9183 | -101.3267
27.922 | -101.3358
27.9257 | -101.3434
27.9294 | -101.3524
27.9331 | -101.3659
27.9367 | -101.3806
27.9404 | -101.3904
27.9441 | -101.3987
27.9478 | -101.4048
27.9515 | -101.404
27.9552 | -101.3998
27.9589 | -101.3943
27.9626 | -101.3874
27.9662 | -101.3818
27.9699 | -101.3769
27.9736 | -101.3712
27.9773 | -101.3664
27.981 | -101.361
27.9847 | -101.3529
27.9884 | -101.345
27.9921 | -101.3362
27.9958 | -101.323
27.9994 | -101.3075
28.0031 | -101.2928