## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 250 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 268
Hood | 02-92B/P | 195
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1014.9 | hPa |  | PRC PRO 01
Humidity | 53 | %/Hr
Internal temperature of calibration bench | 23.22 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1019 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.5 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.92 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 2903 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 18/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 22/06/2018
Technical Manager | Anthony HUE | Approval |  | 28/06/2018
CEO | Robert GLEMOT | Approval |  | 28/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.6066 | 57.5166
0.012209 | -36.1475 | 38.4472
0.018313 | -35.102 | 28.4768
0.024417 | -34.6512 | 22.0836
0.030521 | -34.5433 | 18.2887
0.036626 | -34.4045 | 15.1779
0.04273 | -34.3198 | 12.8807
0.048834 | -34.2675 | 11.3282
0.054938 | -34.1985 | 10.1201
0.061043 | -34.1267 | 9.024
0.067147 | -34.1131 | 8.1239
0.073251 | -34.1093 | 7.3388
0.079355 | -34.1219 | 6.7614
0.08546 | -34.1373 | 6.3015
0.091564 | -34.135 | 5.9392
0.097668 | -34.1044 | 5.6958
0.10377 | -34.0611 | 5.3522
0.10988 | -34.0322 | 4.9962
0.11598 | -34.0215 | 4.639
0.12209 | -34.0337 | 4.3506
0.12819 | -34.0617 | 4.0999
0.13429 | -34.0908 | 3.8576
0.1404 | -34.1093 | 3.6285
0.1465 | -34.1257 | 3.4051
0.15261 | -34.1339 | 3.1932
0.15871 | -34.1313 | 3.007
0.16482 | -34.1169 | 2.8359
0.17092 | -34.0933 | 2.7458
0.17702 | -34.0751 | 2.6315
0.18313 | -34.059 | 2.5376
0.18923 | -34.0473 | 2.4275
0.19534 | -34.0407 | 2.3253
0.20144 | -34.0394 | 2.2149
0.20754 | -34.0413 | 2.1363
0.21365 | -34.0429 | 2.0583
0.21975 | -34.0468 | 1.9885
0.22586 | -34.0543 | 1.9179
0.23196 | -34.0625 | 1.8435
0.23807 | -34.0708 | 1.7629
0.24417 | -34.0732 | 1.6654
0.25027 | -34.077 | 1.5762
0.25638 | -34.0772 | 1.478
0.26248 | -34.0759 | 1.3866
0.26859 | -34.0746 | 1.2975
0.27469 | -34.0749 | 1.2082
0.2808 | -34.0822 | 1.1488
0.2869 | -34.088 | 1.0847
0.293 | -34.0945 | 1.0339
0.29911 | -34.1018 | 1.0029
0.30521 | -34.1083 | 0.97381
0.31132 | -34.1123 | 0.96698
0.31742 | -34.1169 | 0.95203
0.32353 | -34.1205 | 0.93835
0.32963 | -34.1233 | 0.92737
0.33573 | -34.1253 | 0.91555
0.34184 | -34.1279 | 0.89789
0.34794 | -34.1314 | 0.85719
0.35405 | -34.135 | 0.82392
0.36015 | -34.1387 | 0.78621
0.36626 | -34.1414 | 0.74465
0.37236 | -34.1442 | 0.69036
0.37846 | -34.147 | 0.62941
0.38457 | -34.1464 | 0.57033
0.39067 | -34.1474 | 0.51806
0.39678 | -34.1461 | 0.46178
0.40288 | -34.1441 | 0.40861
0.40899 | -34.1405 | 0.35755
0.41509 | -34.1363 | 0.30754
0.42119 | -34.131 | 0.26235
0.4273 | -34.1271 | 0.21485
0.4334 | -34.1233 | 0.16643
0.43951 | -34.12 | 0.11949
0.44561 | -34.118 | 0.07576
0.45172 | -34.1187 | 0.039143
0.45782 | -34.1181 | -0.0011677
0.46392 | -34.1187 | -0.037833
0.47003 | -34.1199 | -0.07158
0.47613 | -34.1215 | -0.10186
0.48224 | -34.1234 | -0.13042
0.48834 | -34.1259 | -0.14956
0.49445 | -34.1278 | -0.17265
0.50055 | -34.1304 | -0.1945
0.50665 | -34.1321 | -0.21049
0.51276 | -34.1334 | -0.23014
0.51886 | -34.1348 | -0.24656
0.52497 | -34.1341 | -0.25456
0.53107 | -34.1346 | -0.26586
0.53717 | -34.135 | -0.27868
0.54328 | -34.1347 | -0.28625
0.54938 | -34.1346 | -0.29618
0.55549 | -34.1339 | -0.3081
0.56159 | -34.1333 | -0.3228
0.5677 | -34.1324 | -0.33591
0.5738 | -34.1314 | -0.35154
0.5799 | -34.1308 | -0.36617
0.58601 | -34.1305 | -0.38485
0.59211 | -34.1306 | -0.40918
0.59822 | -34.1304 | -0.42758
0.60432 | -34.1307 | -0.45161
0.61043 | -34.1308 | -0.47454
0.61653 | -34.1313 | -0.49664
0.62263 | -34.1316 | -0.52099
0.62874 | -34.1325 | -0.54134
0.63484 | -34.1331 | -0.56283
0.64095 | -34.1339 | -0.58455
0.64705 | -34.1347 | -0.6048
0.65316 | -34.1357 | -0.62459
0.65926 | -34.1368 | -0.6468
0.66536 | -34.1368 | -0.6637
0.67147 | -34.1372 | -0.67955
0.67757 | -34.1381 | -0.69728
0.68368 | -34.1383 | -0.71573
0.68978 | -34.1391 | -0.73192
0.69589 | -34.1393 | -0.74889
0.70199 | -34.1392 | -0.76754
0.70809 | -34.1392 | -0.78489
0.7142 | -34.1394 | -0.80179
0.7203 | -34.1393 | -0.81833
0.72641 | -34.1391 | -0.83557
0.73251 | -34.1395 | -0.85391
0.73862 | -34.14 | -0.87506
0.74472 | -34.1408 | -0.89139
0.75082 | -34.1408 | -0.90994
0.75693 | -34.1413 | -0.9309
0.76303 | -34.1417 | -0.95056
0.76914 | -34.1424 | -0.97642
0.77524 | -34.143 | -0.99942
0.78135 | -34.1437 | -1.0209
0.78745 | -34.1447 | -1.0443
0.79355 | -34.1453 | -1.0684
0.79966 | -34.1465 | -1.0934
0.80576 | -34.1478 | -1.1243
0.81187 | -34.1482 | -1.1531
0.81797 | -34.1491 | -1.1838
0.82408 | -34.1499 | -1.2111
0.83018 | -34.1506 | -1.2424
0.83628 | -34.1517 | -1.2731
0.84239 | -34.1516 | -1.3019
0.84849 | -34.1517 | -1.3304
0.8546 | -34.1522 | -1.3581
0.8607 | -34.1523 | -1.3876
0.86681 | -34.1524 | -1.4163
0.87291 | -34.152 | -1.4422
0.87901 | -34.1514 | -1.4651
0.88512 | -34.1513 | -1.4886
0.89122 | -34.1511 | -1.5123
0.89733 | -34.1508 | -1.532
0.90343 | -34.1505 | -1.5539
0.90953 | -34.1505 | -1.5692
0.91564 | -34.1499 | -1.5859
0.92174 | -34.1494 | -1.6006
0.92785 | -34.1488 | -1.6147
0.93395 | -34.1487 | -1.6267
0.94006 | -34.1486 | -1.6406
0.94616 | -34.1484 | -1.6456
0.95226 | -34.1483 | -1.6578
0.95837 | -34.1483 | -1.6699
0.96447 | -34.1487 | -1.6803
0.97058 | -34.1487 | -1.6901
0.97668 | -34.1485 | -1.6981
0.98279 | -34.1486 | -1.7018
0.98889 | -34.1488 | -1.7069
0.99499 | -34.1488 | -1.7156
1.0011 | -34.1492 | -1.7203
1.0072 | -34.1491 | -1.7242
1.0133 | -34.1491 | -1.7283
1.0194 | -34.1493 | -1.7333
1.0255 | -34.149 | -1.7403
1.0316 | -34.1485 | -1.7443
1.0377 | -34.1481 | -1.7499
1.0438 | -34.1475 | -1.7542
1.0499 | -34.1461 | -1.7599
1.056 | -34.1451 | -1.7644
1.0621 | -34.144 | -1.7699
1.0682 | -34.1436 | -1.7746
1.0743 | -34.1427 | -1.7819
1.0805 | -34.1422 | -1.7881
1.0866 | -34.1407 | -1.7984
1.0927 | -34.1395 | -1.8056
1.0988 | -34.1389 | -1.8157
1.1049 | -34.1375 | -1.8228
1.111 | -34.1361 | -1.8335
1.1171 | -34.1348 | -1.8425
1.1232 | -34.1336 | -1.856
1.1293 | -34.1326 | -1.8675
1.1354 | -34.1316 | -1.8793
1.1415 | -34.1306 | -1.8929
1.1476 | -34.1293 | -1.9064
1.1537 | -34.128 | -1.9192
1.1598 | -34.1279 | -1.9373
1.1659 | -34.1274 | -1.9489
1.172 | -34.1267 | -1.962
1.1781 | -34.126 | -1.9755
1.1842 | -34.1255 | -1.9908
1.1903 | -34.1262 | -2.0059
1.1964 | -34.1262 | -2.0203
1.2025 | -34.1261 | -2.0355
1.2086 | -34.1266 | -2.0511
1.2147 | -34.1268 | -2.0661
1.2209 | -34.1275 | -2.0795
1.227 | -34.1293 | -2.0947
1.2331 | -34.1303 | -2.1091
1.2392 | -34.1312 | -2.1248
1.2453 | -34.1324 | -2.137
1.2514 | -34.1337 | -2.1492
1.2575 | -34.1352 | -2.1633
1.2636 | -34.1375 | -2.1749
1.2697 | -34.1392 | -2.1869
1.2758 | -34.1408 | -2.201
1.2819 | -34.1425 | -2.2147
1.288 | -34.1441 | -2.2286
1.2941 | -34.1462 | -2.2394
1.3002 | -34.1482 | -2.2479
1.3063 | -34.1498 | -2.258
1.3124 | -34.1516 | -2.268
1.3185 | -34.1536 | -2.2789
1.3246 | -34.1555 | -2.2898
1.3307 | -34.1576 | -2.2994
1.3368 | -34.1594 | -2.3069
1.3429 | -34.1611 | -2.3157
1.349 | -34.1625 | -2.3241
1.3551 | -34.1643 | -2.3351
1.3613 | -34.1658 | -2.3429
1.3674 | -34.1671 | -2.3511
1.3735 | -34.1684 | -2.3593
1.3796 | -34.1697 | -2.3667
1.3857 | -34.1707 | -2.3738
1.3918 | -34.1718 | -2.3834
1.3979 | -34.1731 | -2.3918
1.404 | -34.1739 | -2.4014
1.4101 | -34.1748 | -2.4101
1.4162 | -34.1757 | -2.4185
1.4223 | -34.1766 | -2.4285
1.4284 | -34.1774 | -2.4365
1.4345 | -34.1781 | -2.4464
1.4406 | -34.1786 | -2.4572
1.4467 | -34.1793 | -2.4662
1.4528 | -34.18 | -2.4752
1.4589 | -34.1808 | -2.4856
1.465 | -34.1811 | -2.4929
1.4711 | -34.1821 | -2.5014
1.4772 | -34.1826 | -2.5132
1.4833 | -34.1832 | -2.5253
1.4894 | -34.184 | -2.5363
1.4955 | -34.1847 | -2.5475
1.5016 | -34.1854 | -2.5592
1.5078 | -34.1858 | -2.5734
1.5139 | -34.1865 | -2.5858
1.52 | -34.1871 | -2.5986
1.5261 | -34.1875 | -2.6098
1.5322 | -34.1877 | -2.622
1.5383 | -34.188 | -2.6339
1.5444 | -34.1887 | -2.6471
1.5505 | -34.1891 | -2.6602
1.5566 | -34.1896 | -2.6729
1.5627 | -34.1902 | -2.6859
1.5688 | -34.1909 | -2.6989
1.5749 | -34.1912 | -2.7111
1.581 | -34.1915 | -2.7278
1.5871 | -34.1921 | -2.7432
1.5932 | -34.1925 | -2.7558
1.5993 | -34.1929 | -2.7723
1.6054 | -34.1932 | -2.7871
1.6115 | -34.1935 | -2.804
1.6176 | -34.1939 | -2.8204
1.6237 | -34.1941 | -2.8347
1.6298 | -34.1944 | -2.8502
1.6359 | -34.1948 | -2.8648
1.642 | -34.1953 | -2.8792
1.6482 | -34.1956 | -2.896
1.6543 | -34.1961 | -2.9127
1.6604 | -34.1965 | -2.9272
1.6665 | -34.1966 | -2.9422
1.6726 | -34.1969 | -2.9586
1.6787 | -34.1974 | -2.9758
1.6848 | -34.1973 | -2.9924
1.6909 | -34.1977 | -3.0087
1.697 | -34.1979 | -3.0251
1.7031 | -34.1979 | -3.0408
1.7092 | -34.1979 | -3.0557
1.7153 | -34.1981 | -3.0718
1.7214 | -34.1977 | -3.0871
1.7275 | -34.1975 | -3.1029
1.7336 | -34.1973 | -3.1191
1.7397 | -34.1974 | -3.1344
1.7458 | -34.1974 | -3.1511
1.7519 | -34.1973 | -3.1673
1.758 | -34.1972 | -3.1823
1.7641 | -34.1973 | -3.1967
1.7702 | -34.1971 | -3.2125
1.7763 | -34.1971 | -3.2263
1.7824 | -34.197 | -3.2426
1.7885 | -34.1967 | -3.2565
1.7947 | -34.1968 | -3.2714
1.8008 | -34.1968 | -3.2844
1.8069 | -34.1967 | -3.298
1.813 | -34.1967 | -3.3122
1.8191 | -34.1964 | -3.3265
1.8252 | -34.1962 | -3.3402
1.8313 | -34.1962 | -3.3547
1.8374 | -34.1962 | -3.3681
1.8435 | -34.1963 | -3.3802
1.8496 | -34.196 | -3.3947
1.8557 | -34.196 | -3.4094
1.8618 | -34.1959 | -3.421
1.8679 | -34.1959 | -3.4339
1.874 | -34.1959 | -3.4471
1.8801 | -34.1958 | -3.46
1.8862 | -34.1959 | -3.4728
1.8923 | -34.1958 | -3.4849
1.8984 | -34.1955 | -3.4955
1.9045 | -34.1954 | -3.5065
1.9106 | -34.1951 | -3.5181
1.9167 | -34.1948 | -3.5293
1.9228 | -34.1948 | -3.5402
1.9289 | -34.1944 | -3.5487
1.9351 | -34.1945 | -3.5576
1.9412 | -34.1946 | -3.5677
1.9473 | -34.1946 | -3.5795
1.9534 | -34.1945 | -3.5905
1.9595 | -34.1943 | -3.602
1.9656 | -34.1938 | -3.6122
1.9717 | -34.1936 | -3.6226
1.9778 | -34.1936 | -3.6338
1.9839 | -34.1936 | -3.645
1.99 | -34.1936 | -3.6548
1.9961 | -34.1933 | -3.6649
2.0022 | -34.1932 | -3.6744
2.0083 | -34.1929 | -3.6842
2.0144 | -34.1928 | -3.694
2.0205 | -34.1929 | -3.7057
2.0266 | -34.193 | -3.7171
2.0327 | -34.193 | -3.7279
2.0388 | -34.1932 | -3.7418
2.0449 | -34.1932 | -3.7519
2.051 | -34.1933 | -3.7619
2.0571 | -34.1931 | -3.7731
2.0632 | -34.1929 | -3.7835
2.0693 | -34.1929 | -3.7946
2.0754 | -34.1931 | -3.8063
2.0816 | -34.1933 | -3.8182
2.0877 | -34.1932 | -3.8292
2.0938 | -34.1934 | -3.8413
2.0999 | -34.1933 | -3.8527
2.106 | -34.1936 | -3.8646
2.1121 | -34.1935 | -3.8767
2.1182 | -34.1937 | -3.8868
2.1243 | -34.1939 | -3.897
2.1304 | -34.1939 | -3.9071
2.1365 | -34.1943 | -3.9185
2.1426 | -34.1944 | -3.9316
2.1487 | -34.1948 | -3.9431
2.1548 | -34.1953 | -3.9561
2.1609 | -34.1958 | -3.9679
2.167 | -34.1962 | -3.9805
2.1731 | -34.1965 | -3.9924
2.1792 | -34.1972 | -4.0037
2.1853 | -34.1977 | -4.0166
2.1914 | -34.1981 | -4.0288
2.1975 | -34.1987 | -4.0407
2.2036 | -34.1992 | -4.0537
2.2097 | -34.1996 | -4.0657
2.2158 | -34.2005 | -4.0759
2.222 | -34.2011 | -4.0859
2.2281 | -34.2014 | -4.0964
2.2342 | -34.2019 | -4.1095
2.2403 | -34.2024 | -4.1194
2.2464 | -34.2031 | -4.1305
2.2525 | -34.2036 | -4.1413
2.2586 | -34.2045 | -4.1511
2.2647 | -34.2053 | -4.1605
2.2708 | -34.2058 | -4.1723
2.2769 | -34.2067 | -4.1839
2.283 | -34.2072 | -4.1932
2.2891 | -34.2077 | -4.2037
2.2952 | -34.2083 | -4.2136
2.3013 | -34.209 | -4.2236
2.3074 | -34.2097 | -4.2332
2.3135 | -34.2101 | -4.2445
2.3196 | -34.2107 | -4.253
2.3257 | -34.2111 | -4.2634
2.3318 | -34.2114 | -4.2722
2.3379 | -34.212 | -4.2836
2.344 | -34.2123 | -4.2945
2.3501 | -34.2127 | -4.3052
2.3562 | -34.2131 | -4.3133
2.3623 | -34.2134 | -4.3221
2.3685 | -34.2136 | -4.3311
2.3746 | -34.2139 | -4.3414
2.3807 | -34.2143 | -4.3501
2.3868 | -34.2145 | -4.3592
2.3929 | -34.2149 | -4.3683
2.399 | -34.2152 | -4.3761
2.4051 | -34.2157 | -4.3855
2.4112 | -34.216 | -4.3943
2.4173 | -34.2163 | -4.4033
2.4234 | -34.2162 | -4.4118
2.4295 | -34.2165 | -4.4202
2.4356 | -34.2166 | -4.4294
2.4417 | -34.2167 | -4.4394
2.4478 | -34.2168 | -4.4491
2.4539 | -34.2171 | -4.4571
2.46 | -34.2167 | -4.4667
2.4661 | -34.2168 | -4.476
2.4722 | -34.217 | -4.4847
2.4783 | -34.2171 | -4.493
2.4844 | -34.2171 | -4.5015
2.4905 | -34.217 | -4.5106
2.4966 | -34.2167 | -4.5194
2.5027 | -34.2166 | -4.5274
2.5089 | -34.2165 | -4.5357
2.515 | -34.2163 | -4.5449
2.5211 | -34.2162 | -4.5533
2.5272 | -34.2158 | -4.5624
2.5333 | -34.2157 | -4.5718
2.5394 | -34.2156 | -4.5808
2.5455 | -34.2153 | -4.59
2.5516 | -34.2153 | -4.5987
2.5577 | -34.2152 | -4.6086
2.5638 | -34.215 | -4.6188
2.5699 | -34.2147 | -4.629
2.576 | -34.2144 | -4.6393
2.5821 | -34.2143 | -4.6492
2.5882 | -34.214 | -4.6588
2.5943 | -34.2139 | -4.6685
2.6004 | -34.2137 | -4.6784
2.6065 | -34.2135 | -4.687
2.6126 | -34.2133 | -4.6961
2.6187 | -34.2132 | -4.7047
2.6248 | -34.213 | -4.7137
2.6309 | -34.2128 | -4.725
2.637 | -34.2129 | -4.7362
2.6431 | -34.2128 | -4.7456
2.6492 | -34.2127 | -4.7555
2.6554 | -34.2126 | -4.7658
2.6615 | -34.2124 | -4.7756
2.6676 | -34.2123 | -4.7861
2.6737 | -34.2124 | -4.7968
2.6798 | -34.2124 | -4.8066
2.6859 | -34.2124 | -4.8171
2.692 | -34.2123 | -4.8267
2.6981 | -34.2123 | -4.8379
2.7042 | -34.2124 | -4.8497
2.7103 | -34.2123 | -4.8598
2.7164 | -34.2125 | -4.8705
2.7225 | -34.2126 | -4.8818
2.7286 | -34.2127 | -4.893
2.7347 | -34.2127 | -4.9028
2.7408 | -34.213 | -4.912
2.7469 | -34.213 | -4.9232
2.753 | -34.2133 | -4.9339
2.7591 | -34.2134 | -4.9458
2.7652 | -34.2134 | -4.9573
2.7713 | -34.2134 | -4.9671
2.7774 | -34.2139 | -4.9784
2.7835 | -34.214 | -4.989
2.7896 | -34.2141 | -4.9981
2.7958 | -34.2145 | -5.0083
2.8019 | -34.215 | -5.018
2.808 | -34.2153 | -5.0299
2.8141 | -34.2158 | -5.0426
2.8202 | -34.2162 | -5.0546
2.8263 | -34.2166 | -5.0654
2.8324 | -34.2169 | -5.0777
2.8385 | -34.2173 | -5.0896
2.8446 | -34.2179 | -5.1012
2.8507 | -34.2183 | -5.1132
2.8568 | -34.2188 | -5.1242
2.8629 | -34.2192 | -5.1355
2.869 | -34.2197 | -5.1466
2.8751 | -34.2202 | -5.1583
2.8812 | -34.2208 | -5.1713
2.8873 | -34.2212 | -5.1824
2.8934 | -34.2216 | -5.1941
2.8995 | -34.222 | -5.2059
2.9056 | -34.2223 | -5.2165
2.9117 | -34.2228 | -5.2277
2.9178 | -34.2235 | -5.2401
2.9239 | -34.224 | -5.2533
2.93 | -34.2245 | -5.2664
2.9361 | -34.2251 | -5.2776
2.9423 | -34.2257 | -5.2894
2.9484 | -34.2263 | -5.3026
2.9545 | -34.227 | -5.3156
2.9606 | -34.2275 | -5.3287
2.9667 | -34.228 | -5.3405
2.9728 | -34.2285 | -5.3524
2.9789 | -34.229 | -5.3649
2.985 | -34.2294 | -5.3784
2.9911 | -34.23 | -5.3909
2.9972 | -34.2305 | -5.4035
3.0033 | -34.2311 | -5.4164
3.0094 | -34.2317 | -5.4292
3.0155 | -34.2323 | -5.4416
3.0216 | -34.2328 | -5.4548
3.0277 | -34.2333 | -5.4682
3.0338 | -34.2339 | -5.4817
3.0399 | -34.2346 | -5.4953
3.046 | -34.2349 | -5.5082
3.0521 | -34.2355 | -5.5205
3.0582 | -34.2358 | -5.5341
3.0643 | -34.2364 | -5.547
3.0704 | -34.2367 | -5.5601
3.0765 | -34.2371 | -5.5721
3.0827 | -34.2374 | -5.5851
3.0888 | -34.2378 | -5.5977
3.0949 | -34.2381 | -5.6105
3.101 | -34.2386 | -5.6244
3.1071 | -34.2389 | -5.6382
3.1132 | -34.2394 | -5.6521
3.1193 | -34.2399 | -5.665
3.1254 | -34.2403 | -5.677
3.1315 | -34.2407 | -5.6911
3.1376 | -34.2409 | -5.7039
3.1437 | -34.2413 | -5.7173
3.1498 | -34.2417 | -5.7314
3.1559 | -34.242 | -5.7445
3.162 | -34.242 | -5.7585
3.1681 | -34.2421 | -5.7725
3.1742 | -34.2425 | -5.786
3.1803 | -34.2429 | -5.7991
3.1864 | -34.2433 | -5.8112
3.1925 | -34.2436 | -5.8247
3.1986 | -34.2437 | -5.8373
3.2047 | -34.2438 | -5.8512
3.2108 | -34.2441 | -5.8628
3.2169 | -34.2445 | -5.8763
3.223 | -34.2444 | -5.8895
3.2292 | -34.2444 | -5.9031
3.2353 | -34.2443 | -5.9125
3.2414 | -34.2445 | -5.9252
3.2475 | -34.2445 | -5.9361
3.2536 | -34.2447 | -5.9497
3.2597 | -34.245 | -5.9608
3.2658 | -34.2451 | -5.9714
3.2719 | -34.2452 | -5.9824
3.278 | -34.2453 | -5.9944
3.2841 | -34.2455 | -6.0063
3.2902 | -34.2455 | -6.0186
3.2963 | -34.2457 | -6.0301
3.3024 | -34.2455 | -6.0402
3.3085 | -34.2458 | -6.0526
3.3146 | -34.2457 | -6.0634
3.3207 | -34.2459 | -6.0755
3.3268 | -34.246 | -6.0876
3.3329 | -34.2461 | -6.0972
3.339 | -34.2461 | -6.1067
3.3451 | -34.2462 | -6.1181
3.3512 | -34.2463 | -6.1275
3.3573 | -34.2464 | -6.1388
3.3634 | -34.2465 | -6.1502
3.3696 | -34.2464 | -6.1619
3.3757 | -34.2464 | -6.1716
3.3818 | -34.2467 | -6.1833
3.3879 | -34.2467 | -6.1937
3.394 | -34.2466 | -6.2041
3.4001 | -34.2466 | -6.2149
3.4062 | -34.2467 | -6.2249
3.4123 | -34.2467 | -6.2346
3.4184 | -34.2467 | -6.2454
3.4245 | -34.2469 | -6.256
3.4306 | -34.2468 | -6.2656
3.4367 | -34.2469 | -6.2753
3.4428 | -34.2472 | -6.2845
3.4489 | -34.2471 | -6.2954
3.455 | -34.247 | -6.3059
3.4611 | -34.2471 | -6.3168
3.4672 | -34.2471 | -6.328
3.4733 | -34.2471 | -6.3382
3.4794 | -34.2471 | -6.3486
3.4855 | -34.2474 | -6.3582
3.4916 | -34.2475 | -6.3679
3.4977 | -34.2474 | -6.3788
3.5038 | -34.2476 | -6.3888
3.5099 | -34.2476 | -6.3983
3.5161 | -34.2478 | -6.4079
3.5222 | -34.2479 | -6.4178
3.5283 | -34.2478 | -6.4279
3.5344 | -34.2478 | -6.4387
3.5405 | -34.248 | -6.4486
3.5466 | -34.2481 | -6.4584
3.5527 | -34.2482 | -6.4691
3.5588 | -34.2483 | -6.479
3.5649 | -34.2483 | -6.4891
3.571 | -34.2484 | -6.4993
3.5771 | -34.2485 | -6.5101
3.5832 | -34.2487 | -6.5196
3.5893 | -34.2489 | -6.5294
3.5954 | -34.249 | -6.5398
3.6015 | -34.2491 | -6.5502
3.6076 | -34.2494 | -6.5599
3.6137 | -34.2494 | -6.5698
3.6198 | -34.2498 | -6.579
3.6259 | -34.2498 | -6.5898
3.632 | -34.2501 | -6.6005
3.6381 | -34.2503 | -6.6104
3.6442 | -34.2505 | -6.6209
3.6503 | -34.2507 | -6.63
3.6565 | -34.251 | -6.6418
3.6626 | -34.2511 | -6.6526
3.6687 | -34.2514 | -6.6632
3.6748 | -34.2518 | -6.6732
3.6809 | -34.2521 | -6.6836
3.687 | -34.2523 | -6.6948
3.6931 | -34.2527 | -6.7048
3.6992 | -34.2529 | -6.7155
3.7053 | -34.2532 | -6.7257
3.7114 | -34.2533 | -6.737
3.7175 | -34.2537 | -6.7469
3.7236 | -34.2539 | -6.7557
3.7297 | -34.254 | -6.7664
3.7358 | -34.2543 | -6.7782
3.7419 | -34.2545 | -6.7899
3.748 | -34.2547 | -6.8005
3.7541 | -34.2548 | -6.8116
3.7602 | -34.2552 | -6.8225
3.7663 | -34.2554 | -6.8327
3.7724 | -34.2558 | -6.8438
3.7785 | -34.2561 | -6.8542
3.7846 | -34.2563 | -6.8649
3.7907 | -34.2568 | -6.876
3.7969 | -34.257 | -6.8875
3.803 | -34.2573 | -6.8976
3.8091 | -34.2577 | -6.908
3.8152 | -34.258 | -6.9192
3.8213 | -34.2582 | -6.931
3.8274 | -34.2586 | -6.9414
3.8335 | -34.259 | -6.9529
3.8396 | -34.2594 | -6.9644
3.8457 | -34.2598 | -6.9747
3.8518 | -34.2603 | -6.9856
3.8579 | -34.2607 | -6.9962
3.864 | -34.261 | -7.007
3.8701 | -34.2615 | -7.0179
3.8762 | -34.2618 | -7.0291
3.8823 | -34.2621 | -7.0397
3.8884 | -34.2625 | -7.0501
3.8945 | -34.263 | -7.0614
3.9006 | -34.2634 | -7.0728
3.9067 | -34.2638 | -7.0846
3.9128 | -34.2642 | -7.0959
3.9189 | -34.2646 | -7.1065
3.925 | -34.2649 | -7.1174
3.9311 | -34.2653 | -7.1273
3.9372 | -34.2657 | -7.1388
3.9434 | -34.2661 | -7.1492
3.9495 | -34.2665 | -7.1607
3.9556 | -34.2669 | -7.1719
3.9617 | -34.2673 | -7.1828
3.9678 | -34.2677 | -7.1935
3.9739 | -34.2681 | -7.2046
3.98 | -34.2685 | -7.2165
3.9861 | -34.2688 | -7.2269
3.9922 | -34.269 | -7.2385
3.9983 | -34.2694 | -7.2486
4.0044 | -34.2698 | -7.2583
4.0105 | -34.2701 | -7.2691
4.0166 | -34.2707 | -7.2801
4.0227 | -34.2711 | -7.2903
4.0288 | -34.2713 | -7.3017
4.0349 | -34.2717 | -7.3115
4.041 | -34.272 | -7.3223
4.0471 | -34.2723 | -7.3325
4.0532 | -34.2727 | -7.3435
4.0593 | -34.273 | -7.355
4.0654 | -34.2734 | -7.3658
4.0715 | -34.2737 | -7.3771
4.0776 | -34.2741 | -7.3874
4.0838 | -34.2745 | -7.3987
4.0899 | -34.2748 | -7.4099
4.096 | -34.2751 | -7.4203
4.1021 | -34.2755 | -7.431
4.1082 | -34.2759 | -7.4413
4.1143 | -34.2762 | -7.4511
4.1204 | -34.2766 | -7.4618
4.1265 | -34.277 | -7.4721
4.1326 | -34.2774 | -7.4823
4.1387 | -34.2779 | -7.4934
4.1448 | -34.2782 | -7.5043
4.1509 | -34.2784 | -7.5149
4.157 | -34.2788 | -7.526
4.1631 | -34.2791 | -7.5372
4.1692 | -34.2795 | -7.5478
4.1753 | -34.2798 | -7.5586
4.1814 | -34.2801 | -7.5683
4.1875 | -34.2804 | -7.5796
4.1936 | -34.2807 | -7.5908
4.1997 | -34.281 | -7.6008
4.2058 | -34.2813 | -7.6122
4.2119 | -34.2816 | -7.6218
4.218 | -34.2818 | -7.6317
4.2241 | -34.2822 | -7.6412
4.2303 | -34.2826 | -7.6523
4.2364 | -34.283 | -7.6637
4.2425 | -34.2832 | -7.6742
4.2486 | -34.2835 | -7.6852
4.2547 | -34.2837 | -7.6942
4.2608 | -34.2839 | -7.704
4.2669 | -34.2842 | -7.7143
4.273 | -34.2845 | -7.725
4.2791 | -34.2847 | -7.7356
4.2852 | -34.2849 | -7.7458
4.2913 | -34.2851 | -7.755
4.2974 | -34.2853 | -7.7658
4.3035 | -34.2856 | -7.7766
4.3096 | -34.2857 | -7.7869
4.3157 | -34.286 | -7.7968
4.3218 | -34.2862 | -7.8068
4.3279 | -34.2864 | -7.817
4.334 | -34.2867 | -7.8268
4.3401 | -34.2869 | -7.8366
4.3462 | -34.2872 | -7.8474
4.3523 | -34.2875 | -7.8569
4.3584 | -34.2877 | -7.8653
4.3645 | -34.2878 | -7.8748
4.3707 | -34.2881 | -7.8841
4.3768 | -34.2883 | -7.8943
4.3829 | -34.2884 | -7.9045
4.389 | -34.2887 | -7.914
4.3951 | -34.289 | -7.9233
4.4012 | -34.2892 | -7.9329
4.4073 | -34.2894 | -7.9429
4.4134 | -34.2896 | -7.9529
4.4195 | -34.2899 | -7.9624
4.4256 | -34.2902 | -7.9719
4.4317 | -34.2904 | -7.9813
4.4378 | -34.2906 | -7.9908
4.4439 | -34.2907 | -8.0002
4.45 | -34.2909 | -8.01
4.4561 | -34.291 | -8.0197
4.4622 | -34.2911 | -8.0291
4.4683 | -34.2913 | -8.0381
4.4744 | -34.2915 | -8.0481
4.4805 | -34.2917 | -8.0571
4.4866 | -34.292 | -8.0662
4.4927 | -34.2921 | -8.0763
4.4988 | -34.2923 | -8.0855
4.5049 | -34.2925 | -8.0945
4.511 | -34.2928 | -8.1034
4.5172 | -34.2931 | -8.1131
4.5233 | -34.2932 | -8.1236
4.5294 | -34.2935 | -8.1328
4.5355 | -34.2937 | -8.1421
4.5416 | -34.294 | -8.1511
4.5477 | -34.2941 | -8.1602
4.5538 | -34.2943 | -8.1689
4.5599 | -34.2947 | -8.1783
4.566 | -34.2948 | -8.1879
4.5721 | -34.295 | -8.197
4.5782 | -34.2953 | -8.2066
4.5843 | -34.2955 | -8.2161
4.5904 | -34.2959 | -8.2252
4.5965 | -34.2961 | -8.2339
4.6026 | -34.2964 | -8.2444
4.6087 | -34.2966 | -8.2532
4.6148 | -34.2968 | -8.2638
4.6209 | -34.2972 | -8.2721
4.627 | -34.2975 | -8.282
4.6331 | -34.2976 | -8.2907
4.6392 | -34.2978 | -8.2998
4.6453 | -34.2981 | -8.3091
4.6514 | -34.2984 | -8.3185
4.6576 | -34.2984 | -8.3288
4.6637 | -34.2986 | -8.3364
4.6698 | -34.2988 | -8.3448
4.6759 | -34.2989 | -8.3545
4.682 | -34.2991 | -8.3635
4.6881 | -34.2993 | -8.3722
4.6942 | -34.2995 | -8.3814
4.7003 | -34.2997 | -8.3907
4.7064 | -34.3 | -8.3998
4.7125 | -34.3003 | -8.4092
4.7186 | -34.3004 | -8.4174
4.7247 | -34.3008 | -8.427
4.7308 | -34.3011 | -8.4361
4.7369 | -34.3012 | -8.4455
4.743 | -34.3014 | -8.4557
4.7491 | -34.3016 | -8.4652
4.7552 | -34.3017 | -8.4743
4.7613 | -34.3019 | -8.4838
4.7674 | -34.3022 | -8.4922
4.7735 | -34.3023 | -8.5021
4.7796 | -34.3025 | -8.5118
4.7857 | -34.3026 | -8.5213
4.7918 | -34.3028 | -8.5298
4.7979 | -34.3032 | -8.5387
4.8041 | -34.3032 | -8.5479
4.8102 | -34.3036 | -8.5569
4.8163 | -34.3038 | -8.5663
4.8224 | -34.3039 | -8.5748
4.8285 | -34.3042 | -8.5837
4.8346 | -34.3045 | -8.5931
4.8407 | -34.3047 | -8.6022
4.8468 | -34.3049 | -8.6118
4.8529 | -34.3051 | -8.6215
4.859 | -34.3053 | -8.6309
4.8651 | -34.3055 | -8.6402
4.8712 | -34.3057 | -8.6495
4.8773 | -34.3059 | -8.6589
4.8834 | -34.3062 | -8.6686
4.8895 | -34.3064 | -8.6785
4.8956 | -34.3066 | -8.6884
4.9017 | -34.3067 | -8.6978
4.9078 | -34.3069 | -8.7076
4.9139 | -34.3072 | -8.7159
4.92 | -34.3074 | -8.7249
4.9261 | -34.3077 | -8.7335
4.9322 | -34.3081 | -8.7427
4.9383 | -34.3084 | -8.7516
4.9445 | -34.3087 | -8.7616
4.9506 | -34.3089 | -8.7704
4.9567 | -34.3091 | -8.7795
4.9628 | -34.3094 | -8.7888
4.9689 | -34.3095 | -8.7985
4.975 | -34.3097 | -8.8076
4.9811 | -34.31 | -8.8173
4.9872 | -34.3103 | -8.8269
4.9933 | -34.3106 | -8.8365
4.9994 | -34.311 | -8.8462
5.0055 | -34.3113 | -8.8558
5.0116 | -34.3116 | -8.8659
5.0177 | -34.312 | -8.8759
5.0238 | -34.3123 | -8.8863
5.0299 | -34.3128 | -8.8972
5.036 | -34.313 | -8.9066
5.0421 | -34.3132 | -8.9162
5.0482 | -34.3135 | -8.9265
5.0543 | -34.3138 | -8.9357
5.0604 | -34.3141 | -8.9459
5.0665 | -34.3145 | -8.9555
5.0726 | -34.3148 | -8.9658
5.0787 | -34.3151 | -8.976
5.0848 | -34.3154 | -8.9859
5.091 | -34.3157 | -8.9957
5.0971 | -34.316 | -9.0057
5.1032 | -34.3164 | -9.016
5.1093 | -34.3167 | -9.0259
5.1154 | -34.3169 | -9.0357
5.1215 | -34.3171 | -9.0457
5.1276 | -34.3175 | -9.0553
5.1337 | -34.3179 | -9.065
5.1398 | -34.3182 | -9.0747
5.1459 | -34.3185 | -9.0851
5.152 | -34.3189 | -9.096
5.1581 | -34.3191 | -9.1065
5.1642 | -34.3193 | -9.1164
5.1703 | -34.3197 | -9.1264
5.1764 | -34.32 | -9.1364
5.1825 | -34.3203 | -9.1466
5.1886 | -34.3206 | -9.1563
5.1947 | -34.3209 | -9.1671
5.2008 | -34.3213 | -9.1768
5.2069 | -34.3216 | -9.188
5.213 | -34.322 | -9.1975
5.2191 | -34.3224 | -9.2078
5.2252 | -34.3226 | -9.218
5.2314 | -34.3229 | -9.2287
5.2375 | -34.3233 | -9.2386
5.2436 | -34.3235 | -9.2487
5.2497 | -34.3239 | -9.2585
5.2558 | -34.3243 | -9.2694
5.2619 | -34.3246 | -9.2801
5.268 | -34.3249 | -9.2905
5.2741 | -34.3253 | -9.3017
5.2802 | -34.3256 | -9.3123
5.2863 | -34.326 | -9.3225
5.2924 | -34.3263 | -9.3336
5.2985 | -34.3267 | -9.3449
5.3046 | -34.3272 | -9.3561
5.3107 | -34.3275 | -9.3668
5.3168 | -34.328 | -9.3773
5.3229 | -34.3283 | -9.3869
5.329 | -34.3287 | -9.3968
5.3351 | -34.3292 | -9.4069
5.3412 | -34.3296 | -9.417
5.3473 | -34.3299 | -9.4279
5.3534 | -34.3303 | -9.4387
5.3595 | -34.3306 | -9.4493
5.3656 | -34.331 | -9.4602
5.3717 | -34.3313 | -9.4706
5.3779 | -34.3317 | -9.4815
5.384 | -34.332 | -9.4921
5.3901 | -34.3324 | -9.5026
5.3962 | -34.3327 | -9.5133
5.4023 | -34.3331 | -9.524
5.4084 | -34.3335 | -9.5343
5.4145 | -34.3338 | -9.5453
5.4206 | -34.3341 | -9.5565
5.4267 | -34.3345 | -9.5677
5.4328 | -34.3349 | -9.5777
5.4389 | -34.3353 | -9.5879
5.445 | -34.3356 | -9.5992
5.4511 | -34.336 | -9.6095
5.4572 | -34.3364 | -9.6207
5.4633 | -34.3367 | -9.631
5.4694 | -34.337 | -9.6414
5.4755 | -34.3374 | -9.6522
5.4816 | -34.3377 | -9.6629
5.4877 | -34.3381 | -9.6734
5.4938 | -34.3384 | -9.6838
5.4999 | -34.3388 | -9.6944
5.506 | -34.3392 | -9.7051
5.5121 | -34.3396 | -9.7162
5.5183 | -34.3399 | -9.7263
5.5244 | -34.3401 | -9.737
5.5305 | -34.3405 | -9.7478
5.5366 | -34.3409 | -9.7581
5.5427 | -34.3412 | -9.7693
5.5488 | -34.3415 | -9.7803
5.5549 | -34.3418 | -9.791
5.561 | -34.3422 | -9.8022
5.5671 | -34.3426 | -9.8124
5.5732 | -34.3429 | -9.8233
5.5793 | -34.3433 | -9.8347
5.5854 | -34.3437 | -9.8455
5.5915 | -34.3439 | -9.8564
5.5976 | -34.3442 | -9.8674
5.6037 | -34.3446 | -9.8779
5.6098 | -34.3449 | -9.8884
5.6159 | -34.3453 | -9.8992
5.622 | -34.3456 | -9.9096
5.6281 | -34.346 | -9.9204
5.6342 | -34.3463 | -9.9312
5.6403 | -34.3466 | -9.9415
5.6464 | -34.347 | -9.9521
5.6525 | -34.3473 | -9.9625
5.6586 | -34.3477 | -9.9724
5.6648 | -34.348 | -9.9836
5.6709 | -34.3484 | -9.9945
5.677 | -34.3487 | -10.006
5.6831 | -34.349 | -10.0162
5.6892 | -34.3494 | -10.0269
5.6953 | -34.3497 | -10.0374
5.7014 | -34.3501 | -10.0486
5.7075 | -34.3504 | -10.0596
5.7136 | -34.3507 | -10.07
5.7197 | -34.3511 | -10.0805
5.7258 | -34.3513 | -10.091
5.7319 | -34.3517 | -10.1017
5.738 | -34.352 | -10.1124
5.7441 | -34.3524 | -10.1229
5.7502 | -34.3528 | -10.1331
5.7563 | -34.3531 | -10.1435
5.7624 | -34.3534 | -10.155
5.7685 | -34.3536 | -10.1654
5.7746 | -34.3539 | -10.1758
5.7807 | -34.3542 | -10.186
5.7868 | -34.3545 | -10.1966
5.7929 | -34.3548 | -10.2072
5.799 | -34.3551 | -10.2177
5.8052 | -34.3554 | -10.2284
5.8113 | -34.3557 | -10.239
5.8174 | -34.356 | -10.2497
5.8235 | -34.3563 | -10.2601
5.8296 | -34.3566 | -10.2713
5.8357 | -34.3568 | -10.282
5.8418 | -34.3572 | -10.2927
5.8479 | -34.3574 | -10.3035
5.854 | -34.3577 | -10.3144
5.8601 | -34.358 | -10.3248
5.8662 | -34.3583 | -10.3356
5.8723 | -34.3585 | -10.3454
5.8784 | -34.3588 | -10.356
5.8845 | -34.3592 | -10.3668
5.8906 | -34.3594 | -10.3774
5.8967 | -34.3597 | -10.3875
5.9028 | -34.3601 | -10.3981
5.9089 | -34.3603 | -10.4081
5.915 | -34.3606 | -10.4186
5.9211 | -34.361 | -10.4292
5.9272 | -34.3612 | -10.4398
5.9333 | -34.3615 | -10.4497
5.9394 | -34.3619 | -10.46
5.9455 | -34.3622 | -10.4701
5.9517 | -34.3626 | -10.4807
5.9578 | -34.3629 | -10.4916
5.9639 | -34.3632 | -10.5017
5.97 | -34.3635 | -10.5123
5.9761 | -34.3639 | -10.5225
5.9822 | -34.3642 | -10.5331
5.9883 | -34.3646 | -10.5434
5.9944 | -34.3648 | -10.554
6.0005 | -34.3652 | -10.5647
6.0066 | -34.3655 | -10.5747
6.0127 | -34.3658 | -10.5845
6.0188 | -34.3661 | -10.5953
6.0249 | -34.3664 | -10.6057
6.031 | -34.3667 | -10.6159
6.0371 | -34.3671 | -10.6265
6.0432 | -34.3674 | -10.6369
6.0493 | -34.3678 | -10.6471
6.0554 | -34.3681 | -10.6575
6.0615 | -34.3684 | -10.6684
6.0676 | -34.3687 | -10.6794
6.0737 | -34.3691 | -10.6899
6.0798 | -34.3693 | -10.7002
6.0859 | -34.3697 | -10.7106
6.0921 | -34.37 | -10.7207
6.0982 | -34.3704 | -10.7314
6.1043 | -34.3708 | -10.7416
6.1104 | -34.3711 | -10.7525
6.1165 | -34.3714 | -10.7621
6.1226 | -34.3718 | -10.7719
6.1287 | -34.3721 | -10.7823
6.1348 | -34.3725 | -10.7927
6.1409 | -34.3728 | -10.8029
6.147 | -34.3731 | -10.813
6.1531 | -34.3734 | -10.8233
6.1592 | -34.3738 | -10.8336
6.1653 | -34.3741 | -10.8438
6.1714 | -34.3744 | -10.8541
6.1775 | -34.3748 | -10.8649
6.1836 | -34.375 | -10.875
6.1897 | -34.3754 | -10.8854
6.1958 | -34.3756 | -10.896
6.2019 | -34.376 | -10.9066
6.208 | -34.3763 | -10.9173
6.2141 | -34.3767 | -10.9279
6.2202 | -34.3771 | -10.9387
6.2263 | -34.3774 | -10.9489
6.2325 | -34.3778 | -10.9588
6.2386 | -34.3782 | -10.9689
6.2447 | -34.3784 | -10.9795
6.2508 | -34.3787 | -10.9897
6.2569 | -34.3791 | -11.0004
6.263 | -34.3795 | -11.0104
6.2691 | -34.3798 | -11.0207
6.2752 | -34.3802 | -11.031
6.2813 | -34.3805 | -11.0414
6.2874 | -34.3808 | -11.0517
6.2935 | -34.3812 | -11.0618
6.2996 | -34.3815 | -11.0722
6.3057 | -34.3819 | -11.0825
6.3118 | -34.3822 | -11.0927
6.3179 | -34.3825 | -11.103
6.324 | -34.3829 | -11.1135
6.3301 | -34.3833 | -11.1236
6.3362 | -34.3836 | -11.1339
6.3423 | -34.384 | -11.1442
6.3484 | -34.3844 | -11.1547
6.3545 | -34.3848 | -11.1644
6.3606 | -34.3852 | -11.1744
6.3667 | -34.3856 | -11.1845
6.3728 | -34.3859 | -11.1948
6.379 | -34.3863 | -11.205
6.3851 | -34.3866 | -11.2153
6.3912 | -34.387 | -11.2253
6.3973 | -34.3874 | -11.2352
6.4034 | -34.3877 | -11.2454
6.4095 | -34.3881 | -11.2559
6.4156 | -34.3884 | -11.2659
6.4217 | -34.3888 | -11.276
6.4278 | -34.3891 | -11.2862
6.4339 | -34.3895 | -11.2964
6.44 | -34.3899 | -11.3063
6.4461 | -34.3902 | -11.317
6.4522 | -34.3906 | -11.3266
6.4583 | -34.3909 | -11.3368
6.4644 | -34.3912 | -11.3474
6.4705 | -34.3916 | -11.3578
6.4766 | -34.392 | -11.3684
6.4827 | -34.3924 | -11.379
6.4888 | -34.3926 | -11.389
6.4949 | -34.393 | -11.3988
6.501 | -34.3934 | -11.4094
6.5071 | -34.3938 | -11.4197
6.5132 | -34.3942 | -11.4297
6.5194 | -34.3945 | -11.4398
6.5255 | -34.3949 | -11.4497
6.5316 | -34.3952 | -11.4601
6.5377 | -34.3956 | -11.4705
6.5438 | -34.396 | -11.4806
6.5499 | -34.3964 | -11.4906
6.556 | -34.3967 | -11.5006
6.5621 | -34.3971 | -11.511
6.5682 | -34.3974 | -11.521
6.5743 | -34.3977 | -11.5319
6.5804 | -34.398 | -11.5421
6.5865 | -34.3983 | -11.5526
6.5926 | -34.3988 | -11.5627
6.5987 | -34.3991 | -11.5726
6.6048 | -34.3995 | -11.5834
6.6109 | -34.3999 | -11.5935
6.617 | -34.4002 | -11.6047
6.6231 | -34.4005 | -11.6151
6.6292 | -34.4009 | -11.6252
6.6353 | -34.4013 | -11.6364
6.6414 | -34.4017 | -11.6469
6.6475 | -34.4021 | -11.6573
6.6536 | -34.4024 | -11.6671
6.6597 | -34.4029 | -11.6771
6.6659 | -34.4032 | -11.6872
6.672 | -34.4036 | -11.6974
6.6781 | -34.404 | -11.7075
6.6842 | -34.4044 | -11.7179
6.6903 | -34.4048 | -11.728
6.6964 | -34.4052 | -11.7383
6.7025 | -34.4055 | -11.7485
6.7086 | -34.4058 | -11.7585
6.7147 | -34.4062 | -11.7689
6.7208 | -34.4065 | -11.7797
6.7269 | -34.4069 | -11.7895
6.733 | -34.4072 | -11.799
6.7391 | -34.4076 | -11.8095
6.7452 | -34.4081 | -11.8191
6.7513 | -34.4085 | -11.8289
6.7574 | -34.4088 | -11.8393
6.7635 | -34.4092 | -11.849
6.7696 | -34.4097 | -11.8593
6.7757 | -34.4101 | -11.8698
6.7818 | -34.4105 | -11.8801
6.7879 | -34.4109 | -11.8904
6.794 | -34.4113 | -11.9004
6.8001 | -34.4116 | -11.9104
6.8063 | -34.412 | -11.9207
6.8124 | -34.4123 | -11.931
6.8185 | -34.4127 | -11.9411
6.8246 | -34.4131 | -11.9513
6.8307 | -34.4134 | -11.9613
6.8368 | -34.4138 | -11.9718
6.8429 | -34.4143 | -11.9818
6.849 | -34.4146 | -11.992
6.8551 | -34.415 | -12.0022
6.8612 | -34.4154 | -12.0121
6.8673 | -34.4158 | -12.022
6.8734 | -34.4161 | -12.033
6.8795 | -34.4166 | -12.0434
6.8856 | -34.4169 | -12.053
6.8917 | -34.4173 | -12.0626
6.8978 | -34.4177 | -12.0727
6.9039 | -34.4181 | -12.0827
6.91 | -34.4185 | -12.0929
6.9161 | -34.4189 | -12.1029
6.9222 | -34.4192 | -12.113
6.9283 | -34.4196 | -12.1228
6.9344 | -34.42 | -12.1327
6.9405 | -34.4204 | -12.1426
6.9466 | -34.4207 | -12.153
6.9528 | -34.4211 | -12.1631
6.9589 | -34.4215 | -12.1731
6.965 | -34.4219 | -12.1831
6.9711 | -34.4223 | -12.1933
6.9772 | -34.4227 | -12.2031
6.9833 | -34.423 | -12.2133
6.9894 | -34.4234 | -12.2239
6.9955 | -34.4238 | -12.2339
7.0016 | -34.4242 | -12.2444
7.0077 | -34.4246 | -12.2549
7.0138 | -34.425 | -12.2653
7.0199 | -34.4253 | -12.2752
7.026 | -34.4256 | -12.2851
7.0321 | -34.426 | -12.2953
7.0382 | -34.4264 | -12.3051
7.0443 | -34.4268 | -12.3155
7.0504 | -34.4271 | -12.3259
7.0565 | -34.4275 | -12.3362
7.0626 | -34.4279 | -12.3461
7.0687 | -34.4282 | -12.3558
7.0748 | -34.4287 | -12.3657
7.0809 | -34.429 | -12.376
7.087 | -34.4294 | -12.3861
7.0932 | -34.4297 | -12.3961
7.0993 | -34.4301 | -12.4063
7.1054 | -34.4305 | -12.416
7.1115 | -34.4308 | -12.426
7.1176 | -34.4311 | -12.4358
7.1237 | -34.4314 | -12.4459
7.1298 | -34.4317 | -12.4554
7.1359 | -34.4321 | -12.4652
7.142 | -34.4324 | -12.4753
7.1481 | -34.4328 | -12.485
7.1542 | -34.4331 | -12.4949
7.1603 | -34.4335 | -12.5048
7.1664 | -34.4339 | -12.5148
7.1725 | -34.4343 | -12.5245
7.1786 | -34.4347 | -12.5346
7.1847 | -34.435 | -12.5444
7.1908 | -34.4354 | -12.5545
7.1969 | -34.4357 | -12.5644
7.203 | -34.4361 | -12.5744
7.2091 | -34.4365 | -12.5842
7.2152 | -34.4368 | -12.5941
7.2213 | -34.4372 | -12.6039
7.2274 | -34.4375 | -12.6138
7.2335 | -34.4379 | -12.6237
7.2397 | -34.4383 | -12.6337
7.2458 | -34.4386 | -12.6431
7.2519 | -34.439 | -12.653
7.258 | -34.4394 | -12.6631
7.2641 | -34.4397 | -12.673
7.2702 | -34.4401 | -12.6832
7.2763 | -34.4405 | -12.6929
7.2824 | -34.4409 | -12.7027
7.2885 | -34.4413 | -12.7126
7.2946 | -34.4417 | -12.7227
7.3007 | -34.442 | -12.7325
7.3068 | -34.4423 | -12.7427
7.3129 | -34.4427 | -12.7523
7.319 | -34.4431 | -12.7623
7.3251 | -34.4434 | -12.7718
7.3312 | -34.4438 | -12.7814
7.3373 | -34.4442 | -12.7909
7.3434 | -34.4447 | -12.8005
7.3495 | -34.445 | -12.8098
7.3556 | -34.4454 | -12.82
7.3617 | -34.4457 | -12.8303
7.3678 | -34.4461 | -12.8402
7.3739 | -34.4465 | -12.8496
7.3801 | -34.4468 | -12.8593
7.3862 | -34.4471 | -12.8691
7.3923 | -34.4475 | -12.8788
7.3984 | -34.4479 | -12.8885
7.4045 | -34.4483 | -12.8982
7.4106 | -34.4487 | -12.9079
7.4167 | -34.449 | -12.9179
7.4228 | -34.4494 | -12.9276
7.4289 | -34.4497 | -12.9371
7.435 | -34.4502 | -12.9466
7.4411 | -34.4505 | -12.956
7.4472 | -34.4509 | -12.9659
7.4533 | -34.4513 | -12.9756
7.4594 | -34.4516 | -12.9856
7.4655 | -34.4521 | -12.9953
7.4716 | -34.4524 | -13.0054
7.4777 | -34.4528 | -13.0146
7.4838 | -34.4531 | -13.0246
7.4899 | -34.4536 | -13.0343
7.496 | -34.454 | -13.0441
7.5021 | -34.4543 | -13.0536
7.5082 | -34.4546 | -13.0634
7.5143 | -34.455 | -13.0731
7.5204 | -34.4554 | -13.0825
7.5266 | -34.4557 | -13.0923
7.5327 | -34.4561 | -13.1019
7.5388 | -34.4564 | -13.1116
7.5449 | -34.4567 | -13.1215
7.551 | -34.4571 | -13.1315
7.5571 | -34.4575 | -13.1411
7.5632 | -34.4578 | -13.1507
7.5693 | -34.4582 | -13.1605
7.5754 | -34.4586 | -13.1703
7.5815 | -34.459 | -13.1796
7.5876 | -34.4593 | -13.1892
7.5937 | -34.4597 | -13.1988
7.5998 | -34.46 | -13.2086
7.6059 | -34.4603 | -13.2182
7.612 | -34.4607 | -13.2278
7.6181 | -34.4611 | -13.2374
7.6242 | -34.4615 | -13.2472
7.6303 | -34.4619 | -13.2567
7.6364 | -34.4622 | -13.2666
7.6425 | -34.4626 | -13.2767
7.6486 | -34.4629 | -13.2864
7.6547 | -34.4633 | -13.2961
7.6608 | -34.4637 | -13.3058
7.667 | -34.4641 | -13.3149
7.6731 | -34.4644 | -13.3248
7.6792 | -34.4647 | -13.3347
7.6853 | -34.465 | -13.3442
7.6914 | -34.4653 | -13.3541
7.6975 | -34.4657 | -13.3639
7.7036 | -34.466 | -13.3737
7.7097 | -34.4664 | -13.3832
7.7158 | -34.4667 | -13.3928
7.7219 | -34.467 | -13.4024
7.728 | -34.4674 | -13.4123
7.7341 | -34.4677 | -13.422
7.7402 | -34.4681 | -13.4316
7.7463 | -34.4684 | -13.4412
7.7524 | -34.4687 | -13.4512
7.7585 | -34.4691 | -13.4611
7.7646 | -34.4694 | -13.4704
7.7707 | -34.4697 | -13.4799
7.7768 | -34.47 | -13.4894
7.7829 | -34.4703 | -13.4992
7.789 | -34.4706 | -13.5085
7.7951 | -34.471 | -13.5184
7.8012 | -34.4713 | -13.5285
7.8073 | -34.4716 | -13.538
7.8135 | -34.4719 | -13.5478
7.8196 | -34.4723 | -13.5573
7.8257 | -34.4727 | -13.5674
7.8318 | -34.473 | -13.5774
7.8379 | -34.4734 | -13.5866
7.844 | -34.4736 | -13.5958
7.8501 | -34.4739 | -13.6057
7.8562 | -34.4742 | -13.616
7.8623 | -34.4746 | -13.6257
7.8684 | -34.4749 | -13.6351
7.8745 | -34.4752 | -13.6446
7.8806 | -34.4755 | -13.6538
7.8867 | -34.4759 | -13.6633
7.8928 | -34.4762 | -13.6729
7.8989 | -34.4766 | -13.6821
7.905 | -34.477 | -13.6921
7.9111 | -34.4773 | -13.7018
7.9172 | -34.4776 | -13.7112
7.9233 | -34.4779 | -13.721
7.9294 | -34.4783 | -13.7308
7.9355 | -34.4786 | -13.74
7.9416 | -34.4789 | -13.7496
7.9477 | -34.4792 | -13.7588
7.9539 | -34.4796 | -13.768
7.96 | -34.4799 | -13.7779
7.9661 | -34.4802 | -13.7876
7.9722 | -34.4806 | -13.7974
7.9783 | -34.4809 | -13.807
7.9844 | -34.4813 | -13.817
7.9905 | -34.4816 | -13.8266
7.9966 | -34.4818 | -13.8361
8.0027 | -34.4822 | -13.8461
8.0088 | -34.4825 | -13.8553
8.0149 | -34.4828 | -13.8646
8.021 | -34.4831 | -13.8743
8.0271 | -34.4835 | -13.8843
8.0332 | -34.4837 | -13.8939
8.0393 | -34.4841 | -13.9037
8.0454 | -34.4844 | -13.9136
8.0515 | -34.4848 | -13.9231
8.0576 | -34.485 | -13.933
8.0637 | -34.4853 | -13.9422
8.0698 | -34.4857 | -13.9513
8.0759 | -34.486 | -13.9606
8.082 | -34.4863 | -13.9704
8.0881 | -34.4866 | -13.9798
8.0942 | -34.4869 | -13.9899
8.1004 | -34.4872 | -13.9999
8.1065 | -34.4876 | -14.0094
8.1126 | -34.4879 | -14.019
8.1187 | -34.4883 | -14.0287
8.1248 | -34.4886 | -14.0385
8.1309 | -34.4889 | -14.0481
8.137 | -34.4892 | -14.0577
8.1431 | -34.4896 | -14.0675
8.1492 | -34.4899 | -14.0772
8.1553 | -34.4902 | -14.0868
8.1614 | -34.4905 | -14.0967
8.1675 | -34.4908 | -14.1061
8.1736 | -34.4911 | -14.1161
8.1797 | -34.4914 | -14.1259
8.1858 | -34.4917 | -14.1355
8.1919 | -34.492 | -14.1451
8.198 | -34.4923 | -14.1544
8.2041 | -34.4927 | -14.1643
8.2102 | -34.493 | -14.174
8.2163 | -34.4934 | -14.1836
8.2224 | -34.4937 | -14.1935
8.2285 | -34.494 | -14.203
8.2346 | -34.4943 | -14.2125
8.2408 | -34.4947 | -14.2221
8.2469 | -34.495 | -14.2316
8.253 | -34.4953 | -14.2412
8.2591 | -34.4956 | -14.2506
8.2652 | -34.496 | -14.2601
8.2713 | -34.4964 | -14.2698
8.2774 | -34.4967 | -14.2794
8.2835 | -34.497 | -14.2895
8.2896 | -34.4973 | -14.2986
8.2957 | -34.4977 | -14.3082
8.3018 | -34.498 | -14.3179
8.3079 | -34.4983 | -14.3274
8.314 | -34.4986 | -14.3369
8.3201 | -34.499 | -14.3465
8.3262 | -34.4993 | -14.3561
8.3323 | -34.4996 | -14.3657
8.3384 | -34.5 | -14.3758
8.3445 | -34.5003 | -14.385
8.3506 | -34.5006 | -14.3949
8.3567 | -34.501 | -14.4048
8.3628 | -34.5013 | -14.4143
8.3689 | -34.5016 | -14.424
8.375 | -34.5019 | -14.4336
8.3812 | -34.5022 | -14.4432
8.3873 | -34.5026 | -14.4529
8.3934 | -34.5029 | -14.4625
8.3995 | -34.5032 | -14.4723
8.4056 | -34.5036 | -14.4821
8.4117 | -34.5039 | -14.4918
8.4178 | -34.5042 | -14.5011
8.4239 | -34.5045 | -14.5104
8.43 | -34.5049 | -14.52
8.4361 | -34.5052 | -14.53
8.4422 | -34.5055 | -14.5396
8.4483 | -34.5058 | -14.5492
8.4544 | -34.5062 | -14.559
8.4605 | -34.5065 | -14.5686
8.4666 | -34.5067 | -14.5784
8.4727 | -34.507 | -14.5879
8.4788 | -34.5074 | -14.5974
8.4849 | -34.5077 | -14.6074
8.491 | -34.508 | -14.6169
8.4971 | -34.5083 | -14.6266
8.5032 | -34.5087 | -14.6363
8.5093 | -34.509 | -14.6459
8.5154 | -34.5093 | -14.6553
8.5215 | -34.5096 | -14.665
8.5277 | -34.5099 | -14.6747
8.5338 | -34.5102 | -14.6841
8.5399 | -34.5106 | -14.6941
8.546 | -34.5109 | -14.7034
8.5521 | -34.5112 | -14.7132
8.5582 | -34.5115 | -14.723
8.5643 | -34.5119 | -14.7325
8.5704 | -34.5122 | -14.7418
8.5765 | -34.5125 | -14.7513
8.5826 | -34.5129 | -14.7609
8.5887 | -34.5132 | -14.7705
8.5948 | -34.5135 | -14.7798
8.6009 | -34.5139 | -14.7893
8.607 | -34.5141 | -14.7992
8.6131 | -34.5144 | -14.8091
8.6192 | -34.5148 | -14.8184
8.6253 | -34.5151 | -14.8282
8.6314 | -34.5155 | -14.8378
8.6375 | -34.5158 | -14.8477
8.6436 | -34.5162 | -14.8573
8.6497 | -34.5165 | -14.867
8.6558 | -34.5168 | -14.8764
8.6619 | -34.5171 | -14.8862
8.6681 | -34.5174 | -14.8958
8.6742 | -34.5178 | -14.9056
8.6803 | -34.5181 | -14.9155
8.6864 | -34.5185 | -14.9254
8.6925 | -34.5189 | -14.935
8.6986 | -34.5192 | -14.945
8.7047 | -34.5195 | -14.9548
8.7108 | -34.5198 | -14.9645
8.7169 | -34.5202 | -14.9739
8.723 | -34.5205 | -14.9835
8.7291 | -34.5208 | -14.993
8.7352 | -34.5212 | -15.003
8.7413 | -34.5215 | -15.0125
8.7474 | -34.5219 | -15.0219
8.7535 | -34.5223 | -15.0317
8.7596 | -34.5226 | -15.0419
8.7657 | -34.5229 | -15.0515
8.7718 | -34.5233 | -15.0609
8.7779 | -34.5236 | -15.071
8.784 | -34.5239 | -15.0805
8.7901 | -34.5241 | -15.0906
8.7962 | -34.5245 | -15.1003
8.8023 | -34.5249 | -15.1096
8.8084 | -34.5252 | -15.1195
8.8146 | -34.5256 | -15.129
8.8207 | -34.5259 | -15.1387
8.8268 | -34.5263 | -15.1482
8.8329 | -34.5266 | -15.158
8.839 | -34.527 | -15.1677
8.8451 | -34.5273 | -15.1771
8.8512 | -34.5277 | -15.1868
8.8573 | -34.528 | -15.1963
8.8634 | -34.5284 | -15.206
8.8695 | -34.5287 | -15.2157
8.8756 | -34.5291 | -15.2251
8.8817 | -34.5294 | -15.235
8.8878 | -34.5297 | -15.2449
8.8939 | -34.53 | -15.2543
8.9 | -34.5303 | -15.2637
8.9061 | -34.5307 | -15.2734
8.9122 | -34.5311 | -15.2829
8.9183 | -34.5314 | -15.2924
8.9244 | -34.5317 | -15.3018
8.9305 | -34.5321 | -15.3115
8.9366 | -34.5325 | -15.3209
8.9427 | -34.5329 | -15.3307
8.9488 | -34.5332 | -15.3403
8.955 | -34.5336 | -15.3499
8.9611 | -34.5339 | -15.3597
8.9672 | -34.5343 | -15.3697
8.9733 | -34.5347 | -15.3792
8.9794 | -34.535 | -15.389
8.9855 | -34.5354 | -15.3985
8.9916 | -34.5357 | -15.408
8.9977 | -34.5361 | -15.4176
9.0038 | -34.5365 | -15.4274
9.0099 | -34.5368 | -15.4371
9.016 | -34.5372 | -15.4468
9.0221 | -34.5376 | -15.4568
9.0282 | -34.538 | -15.4666
9.0343 | -34.5383 | -15.4765
9.0404 | -34.5388 | -15.4861
9.0465 | -34.5391 | -15.4961
9.0526 | -34.5395 | -15.5056
9.0587 | -34.5399 | -15.5152
9.0648 | -34.5403 | -15.5248
9.0709 | -34.5406 | -15.5342
9.077 | -34.541 | -15.5445
9.0831 | -34.5413 | -15.5541
9.0892 | -34.5416 | -15.5638
9.0953 | -34.542 | -15.5736
9.1015 | -34.5423 | -15.5832
9.1076 | -34.5427 | -15.593
9.1137 | -34.5431 | -15.6026
9.1198 | -34.5435 | -15.6118
9.1259 | -34.5439 | -15.6215
9.132 | -34.5443 | -15.6315
9.1381 | -34.5447 | -15.641
9.1442 | -34.5452 | -15.6503
9.1503 | -34.5455 | -15.66
9.1564 | -34.5459 | -15.6697
9.1625 | -34.5462 | -15.6794
9.1686 | -34.5466 | -15.689
9.1747 | -34.547 | -15.6988
9.1808 | -34.5474 | -15.7087
9.1869 | -34.5477 | -15.7187
9.193 | -34.5481 | -15.7285
9.1991 | -34.5484 | -15.7379
9.2052 | -34.5488 | -15.7474
9.2113 | -34.5492 | -15.7573
9.2174 | -34.5496 | -15.7676
9.2235 | -34.55 | -15.777
9.2296 | -34.5505 | -15.7868
9.2357 | -34.5508 | -15.7965
9.2419 | -34.5512 | -15.8061
9.248 | -34.5516 | -15.816
9.2541 | -34.552 | -15.8259
9.2602 | -34.5524 | -15.8357
9.2663 | -34.5527 | -15.8452
9.2724 | -34.5531 | -15.8549
9.2785 | -34.5535 | -15.8646
9.2846 | -34.5539 | -15.8744
9.2907 | -34.5543 | -15.884
9.2968 | -34.5546 | -15.8938
9.3029 | -34.555 | -15.9036
9.309 | -34.5554 | -15.9132
9.3151 | -34.5558 | -15.9228
9.3212 | -34.5562 | -15.9326
9.3273 | -34.5566 | -15.9422
9.3334 | -34.5571 | -15.9517
9.3395 | -34.5575 | -15.9612
9.3456 | -34.5579 | -15.9707
9.3517 | -34.5583 | -15.98
9.3578 | -34.5587 | -15.9895
9.3639 | -34.5592 | -15.9993
9.37 | -34.5596 | -16.009
9.3761 | -34.56 | -16.0186
9.3822 | -34.5603 | -16.0282
9.3884 | -34.5608 | -16.0378
9.3945 | -34.5612 | -16.0475
9.4006 | -34.5615 | -16.0572
9.4067 | -34.5619 | -16.067
9.4128 | -34.5623 | -16.0766
9.4189 | -34.5628 | -16.0864
9.425 | -34.5632 | -16.0959
9.4311 | -34.5636 | -16.1055
9.4372 | -34.564 | -16.1152
9.4433 | -34.5644 | -16.1249
9.4494 | -34.5648 | -16.1346
9.4555 | -34.5652 | -16.1442
9.4616 | -34.5656 | -16.1538
9.4677 | -34.5661 | -16.1633
9.4738 | -34.5665 | -16.1732
9.4799 | -34.5668 | -16.1828
9.486 | -34.5671 | -16.1928
9.4921 | -34.5676 | -16.2029
9.4982 | -34.568 | -16.2121
9.5043 | -34.5684 | -16.2215
9.5104 | -34.5688 | -16.231
9.5165 | -34.5693 | -16.2406
9.5226 | -34.5697 | -16.2505
9.5288 | -34.5701 | -16.2601
9.5349 | -34.5705 | -16.2698
9.541 | -34.5709 | -16.2794
9.5471 | -34.5713 | -16.2891
9.5532 | -34.5717 | -16.2987
9.5593 | -34.5722 | -16.3083
9.5654 | -34.5726 | -16.3181
9.5715 | -34.573 | -16.328
9.5776 | -34.5734 | -16.3379
9.5837 | -34.5738 | -16.3472
9.5898 | -34.5742 | -16.3571
9.5959 | -34.5745 | -16.3666
9.602 | -34.5749 | -16.3762
9.6081 | -34.5753 | -16.3861
9.6142 | -34.5758 | -16.3956
9.6203 | -34.5762 | -16.4053
9.6264 | -34.5765 | -16.4146
9.6325 | -34.5769 | -16.4244
9.6386 | -34.5773 | -16.434
9.6447 | -34.5777 | -16.4436
9.6508 | -34.5781 | -16.4533
9.6569 | -34.5785 | -16.4629
9.663 | -34.579 | -16.4726
9.6691 | -34.5794 | -16.4822
9.6753 | -34.5799 | -16.4919
9.6814 | -34.5803 | -16.5013
9.6875 | -34.5807 | -16.5108
9.6936 | -34.5812 | -16.5202
9.6997 | -34.5816 | -16.5301
9.7058 | -34.582 | -16.5397
9.7119 | -34.5824 | -16.5493
9.718 | -34.5828 | -16.5592
9.7241 | -34.5832 | -16.5688
9.7302 | -34.5836 | -16.5785
9.7363 | -34.584 | -16.5879
9.7424 | -34.5845 | -16.5975
9.7485 | -34.5849 | -16.607
9.7546 | -34.5853 | -16.6163
9.7607 | -34.5857 | -16.6262
9.7668 | -34.5861 | -16.6362
9.7729 | -34.5866 | -16.6456
9.779 | -34.5871 | -16.6552
9.7851 | -34.5874 | -16.6649
9.7912 | -34.5878 | -16.6751
9.7973 | -34.5882 | -16.6849
9.8034 | -34.5886 | -16.6947
9.8095 | -34.589 | -16.7042
9.8157 | -34.5895 | -16.7138
9.8218 | -34.5899 | -16.7231
9.8279 | -34.5903 | -16.7328
9.834 | -34.5907 | -16.7423
9.8401 | -34.5912 | -16.7518
9.8462 | -34.5916 | -16.7611
9.8523 | -34.5921 | -16.7706
9.8584 | -34.5925 | -16.7805
9.8645 | -34.5928 | -16.7901
9.8706 | -34.5932 | -16.7996
9.8767 | -34.5936 | -16.8088
9.8828 | -34.5941 | -16.8179
9.8889 | -34.5946 | -16.8278
9.895 | -34.5949 | -16.837
9.9011 | -34.5953 | -16.8462
9.9072 | -34.5958 | -16.8554
9.9133 | -34.5962 | -16.8648
9.9194 | -34.5967 | -16.8741
9.9255 | -34.597 | -16.8838
9.9316 | -34.5975 | -16.8931
9.9377 | -34.598 | -16.9025
9.9438 | -34.5984 | -16.9119
9.9499 | -34.5988 | -16.9216
9.956 | -34.5992 | -16.9309
9.9622 | -34.5997 | -16.9403
9.9683 | -34.6001 | -16.9499
9.9744 | -34.6005 | -16.9593
9.9805 | -34.6009 | -16.9689
9.9866 | -34.6013 | -16.9781
9.9927 | -34.6018 | -16.9878
9.9988 | -34.6022 | -16.9971
10.0049 | -34.6026 | -17.0065
10.011 | -34.6031 | -17.016
10.0171 | -34.6035 | -17.0255
10.0232 | -34.604 | -17.0347
10.0293 | -34.6044 | -17.044
10.0354 | -34.6049 | -17.0532
10.0415 | -34.6053 | -17.0626
10.0476 | -34.6058 | -17.0724
10.0537 | -34.6061 | -17.0822
10.0598 | -34.6066 | -17.0918
10.0659 | -34.607 | -17.1013
10.072 | -34.6074 | -17.1105
10.0781 | -34.6079 | -17.12
10.0842 | -34.6083 | -17.1296
10.0903 | -34.6088 | -17.1393
10.0964 | -34.6093 | -17.1485
10.1026 | -34.6097 | -17.1579
10.1087 | -34.6101 | -17.1674
10.1148 | -34.6106 | -17.177
10.1209 | -34.611 | -17.1862
10.127 | -34.6114 | -17.1954
10.1331 | -34.6119 | -17.2047
10.1392 | -34.6123 | -17.214
10.1453 | -34.6128 | -17.2235
10.1514 | -34.6132 | -17.2326
10.1575 | -34.6136 | -17.2419
10.1636 | -34.6141 | -17.2506
10.1697 | -34.6145 | -17.26
10.1758 | -34.615 | -17.2692
10.1819 | -34.6154 | -17.2787
10.188 | -34.6158 | -17.2879
10.1941 | -34.6162 | -17.2973
10.2002 | -34.6166 | -17.3061
10.2063 | -34.617 | -17.3157
10.2124 | -34.6174 | -17.3252
10.2185 | -34.6179 | -17.3345
10.2246 | -34.6183 | -17.3438
10.2307 | -34.6187 | -17.3533
10.2368 | -34.6192 | -17.3618
10.2429 | -34.6197 | -17.3709
10.2491 | -34.6201 | -17.38
10.2552 | -34.6205 | -17.3897
10.2613 | -34.621 | -17.3993
10.2674 | -34.6213 | -17.4089
10.2735 | -34.6217 | -17.4184
10.2796 | -34.6222 | -17.4278
10.2857 | -34.6226 | -17.4372
10.2918 | -34.623 | -17.4467
10.2979 | -34.6235 | -17.4562
10.304 | -34.624 | -17.4658
10.3101 | -34.6244 | -17.4751
10.3162 | -34.6248 | -17.4844
10.3223 | -34.6253 | -17.4938
10.3284 | -34.6257 | -17.5031
10.3345 | -34.6261 | -17.5122
10.3406 | -34.6266 | -17.5214
10.3467 | -34.627 | -17.5307
10.3528 | -34.6274 | -17.5398
10.3589 | -34.6278 | -17.5493
10.365 | -34.6283 | -17.5588
10.3711 | -34.6286 | -17.5683
10.3772 | -34.629 | -17.5777
10.3833 | -34.6295 | -17.587
10.3895 | -34.6299 | -17.5961
10.3956 | -34.6303 | -17.6052
10.4017 | -34.6307 | -17.6144
10.4078 | -34.6312 | -17.6238
10.4139 | -34.6316 | -17.6331
10.42 | -34.6321 | -17.6425
10.4261 | -34.6326 | -17.6516
10.4322 | -34.633 | -17.6608
10.4383 | -34.6334 | -17.67
10.4444 | -34.6338 | -17.679
10.4505 | -34.6343 | -17.6883
10.4566 | -34.6347 | -17.6975
10.4627 | -34.6351 | -17.7068
10.4688 | -34.6356 | -17.7156
10.4749 | -34.636 | -17.7247
10.481 | -34.6365 | -17.7339
10.4871 | -34.6369 | -17.7433
10.4932 | -34.6374 | -17.7525
10.4993 | -34.6378 | -17.7616
10.5054 | -34.6383 | -17.7709
10.5115 | -34.6387 | -17.7802
10.5176 | -34.6391 | -17.7891
10.5237 | -34.6396 | -17.7983
10.5298 | -34.64 | -17.8072
10.536 | -34.6405 | -17.8166
10.5421 | -34.6409 | -17.8254
10.5482 | -34.6413 | -17.835
10.5543 | -34.6418 | -17.8441
10.5604 | -34.6423 | -17.853
10.5665 | -34.6427 | -17.862
10.5726 | -34.6431 | -17.871
10.5787 | -34.6436 | -17.8803
10.5848 | -34.6441 | -17.8891
10.5909 | -34.6445 | -17.8983
10.597 | -34.6449 | -17.9075
10.6031 | -34.6454 | -17.9169
10.6092 | -34.6458 | -17.9263
10.6153 | -34.6462 | -17.9358
10.6214 | -34.6467 | -17.9448
10.6275 | -34.6471 | -17.9538
10.6336 | -34.6476 | -17.9634
10.6397 | -34.648 | -17.9724
10.6458 | -34.6484 | -17.9814
10.6519 | -34.6488 | -17.9904
10.658 | -34.6492 | -17.9993
10.6641 | -34.6496 | -18.0085
10.6702 | -34.6501 | -18.0179
10.6764 | -34.6506 | -18.027
10.6825 | -34.6511 | -18.0363
10.6886 | -34.6515 | -18.0455
10.6947 | -34.6519 | -18.0544
10.7008 | -34.6524 | -18.0635
10.7069 | -34.6529 | -18.0727
10.713 | -34.6533 | -18.0817
10.7191 | -34.6538 | -18.0909
10.7252 | -34.6542 | -18.0997
10.7313 | -34.6547 | -18.1089
10.7374 | -34.6551 | -18.1178
10.7435 | -34.6555 | -18.1272
10.7496 | -34.656 | -18.1361
10.7557 | -34.6564 | -18.1453
10.7618 | -34.6569 | -18.1543
10.7679 | -34.6573 | -18.1634
10.774 | -34.6577 | -18.1725
10.7801 | -34.6582 | -18.182
10.7862 | -34.6586 | -18.191
10.7923 | -34.6591 | -18.2002
10.7984 | -34.6595 | -18.2093
10.8045 | -34.66 | -18.2185
10.8106 | -34.6604 | -18.2276
10.8168 | -34.6608 | -18.2367
10.8229 | -34.6613 | -18.2458
10.829 | -34.6617 | -18.255
10.8351 | -34.6622 | -18.2641
10.8412 | -34.6626 | -18.2731
10.8473 | -34.663 | -18.2819
10.8534 | -34.6635 | -18.2912
10.8595 | -34.6639 | -18.3001
10.8656 | -34.6644 | -18.309
10.8717 | -34.6649 | -18.3176
10.8778 | -34.6653 | -18.3266
10.8839 | -34.6658 | -18.3356
10.89 | -34.6662 | -18.3444
10.8961 | -34.6666 | -18.3535
10.9022 | -34.6671 | -18.3625
10.9083 | -34.6676 | -18.3718
10.9144 | -34.668 | -18.3807
10.9205 | -34.6685 | -18.3898
10.9266 | -34.669 | -18.3991
10.9327 | -34.6695 | -18.4081
10.9388 | -34.6699 | -18.4171
10.9449 | -34.6703 | -18.4262
10.951 | -34.6708 | -18.4352
10.9571 | -34.6712 | -18.4443
10.9633 | -34.6717 | -18.4537
10.9694 | -34.6721 | -18.4631
10.9755 | -34.6726 | -18.4721
10.9816 | -34.6731 | -18.4812
10.9877 | -34.6735 | -18.4902
10.9938 | -34.674 | -18.4995
10.9999 | -34.6744 | -18.5085
11.006 | -34.6749 | -18.5174
11.0121 | -34.6753 | -18.5268
11.0182 | -34.6758 | -18.5355
11.0243 | -34.6763 | -18.5445
11.0304 | -34.6767 | -18.5534
11.0365 | -34.6771 | -18.5625
11.0426 | -34.6776 | -18.5713
11.0487 | -34.6781 | -18.5805
11.0548 | -34.6785 | -18.5894
11.0609 | -34.679 | -18.5987
11.067 | -34.6794 | -18.6079
11.0731 | -34.6799 | -18.617
11.0792 | -34.6803 | -18.6261
11.0853 | -34.6807 | -18.6349
11.0914 | -34.6812 | -18.6438
11.0975 | -34.6817 | -18.6531
11.1037 | -34.6821 | -18.6617
11.1098 | -34.6825 | -18.6709
11.1159 | -34.6831 | -18.6801
11.122 | -34.6835 | -18.6894
11.1281 | -34.6839 | -18.6982
11.1342 | -34.6844 | -18.707
11.1403 | -34.6848 | -18.7164
11.1464 | -34.6853 | -18.7253
11.1525 | -34.6857 | -18.7344
11.1586 | -34.6862 | -18.744
11.1647 | -34.6867 | -18.7528
11.1708 | -34.6871 | -18.7617
11.1769 | -34.6875 | -18.7701
11.183 | -34.688 | -18.7789
11.1891 | -34.6885 | -18.7883
11.1952 | -34.6889 | -18.7976
11.2013 | -34.6893 | -18.8065
11.2074 | -34.6898 | -18.8152
11.2135 | -34.6901 | -18.8242
11.2196 | -34.6905 | -18.8332
11.2257 | -34.6909 | -18.8421
11.2318 | -34.6913 | -18.8513
11.2379 | -34.6917 | -18.8602
11.244 | -34.6922 | -18.8692
11.2502 | -34.6926 | -18.878
11.2563 | -34.6931 | -18.8871
11.2624 | -34.6935 | -18.8961
11.2685 | -34.694 | -18.9048
11.2746 | -34.6944 | -18.9135
11.2807 | -34.6948 | -18.9229
11.2868 | -34.6952 | -18.9322
11.2929 | -34.6957 | -18.9409
11.299 | -34.6962 | -18.9502
11.3051 | -34.6966 | -18.9592
11.3112 | -34.697 | -18.9682
11.3173 | -34.6975 | -18.9773
11.3234 | -34.6979 | -18.9863
11.3295 | -34.6984 | -18.9952
11.3356 | -34.6989 | -19.004
11.3417 | -34.6993 | -19.013
11.3478 | -34.6997 | -19.0219
11.3539 | -34.7001 | -19.0307
11.36 | -34.7006 | -19.0398
11.3661 | -34.7011 | -19.0488
11.3722 | -34.7015 | -19.0577
11.3783 | -34.702 | -19.0665
11.3844 | -34.7024 | -19.0756
11.3906 | -34.7028 | -19.0848
11.3967 | -34.7032 | -19.0938
11.4028 | -34.7036 | -19.1029
11.4089 | -34.7041 | -19.1117
11.415 | -34.7045 | -19.1206
11.4211 | -34.7049 | -19.1294
11.4272 | -34.7053 | -19.1384
11.4333 | -34.7058 | -19.1475
11.4394 | -34.7062 | -19.1566
11.4455 | -34.7066 | -19.1658
11.4516 | -34.707 | -19.1746
11.4577 | -34.7075 | -19.184
11.4638 | -34.708 | -19.1925
11.4699 | -34.7083 | -19.2019
11.476 | -34.7088 | -19.2113
11.4821 | -34.7092 | -19.2201
11.4882 | -34.7097 | -19.2293
11.4943 | -34.7101 | -19.2381
11.5004 | -34.7105 | -19.247
11.5065 | -34.711 | -19.2559
11.5126 | -34.7115 | -19.2646
11.5187 | -34.7119 | -19.2735
11.5248 | -34.7124 | -19.2824
11.5309 | -34.7128 | -19.2913
11.5371 | -34.7132 | -19.3001
11.5432 | -34.7136 | -19.3093
11.5493 | -34.714 | -19.3184
11.5554 | -34.7145 | -19.3271
11.5615 | -34.7149 | -19.3363
11.5676 | -34.7154 | -19.3455
11.5737 | -34.7159 | -19.3541
11.5798 | -34.7163 | -19.3629
11.5859 | -34.7168 | -19.3717
11.592 | -34.7172 | -19.3806
11.5981 | -34.7176 | -19.3895
11.6042 | -34.7181 | -19.3985
11.6103 | -34.7185 | -19.4074
11.6164 | -34.7189 | -19.4162
11.6225 | -34.7194 | -19.4251
11.6286 | -34.7199 | -19.4342
11.6347 | -34.7203 | -19.4435
11.6408 | -34.7208 | -19.4523
11.6469 | -34.7212 | -19.4615
11.653 | -34.7217 | -19.4701
11.6591 | -34.7221 | -19.479
11.6652 | -34.7226 | -19.4881
11.6713 | -34.723 | -19.497
11.6775 | -34.7235 | -19.5058
11.6836 | -34.7239 | -19.5148
11.6897 | -34.7244 | -19.5236
11.6958 | -34.7248 | -19.5328
11.7019 | -34.7253 | -19.5419
11.708 | -34.7257 | -19.551
11.7141 | -34.7263 | -19.5601
11.7202 | -34.7267 | -19.5688
11.7263 | -34.7271 | -19.5778
11.7324 | -34.7276 | -19.5869
11.7385 | -34.728 | -19.5959
11.7446 | -34.7284 | -19.6051
11.7507 | -34.7289 | -19.6139
11.7568 | -34.7293 | -19.623
11.7629 | -34.7297 | -19.632
11.769 | -34.7302 | -19.6409
11.7751 | -34.7307 | -19.6496
11.7812 | -34.7311 | -19.6584
11.7873 | -34.7315 | -19.6676
11.7934 | -34.732 | -19.6765
11.7995 | -34.7325 | -19.6851
11.8056 | -34.7329 | -19.6942
11.8117 | -34.7332 | -19.7032
11.8178 | -34.7337 | -19.712
11.824 | -34.7342 | -19.7207
11.8301 | -34.7347 | -19.7291
11.8362 | -34.7351 | -19.7377
11.8423 | -34.7355 | -19.7467
11.8484 | -34.736 | -19.7555
11.8545 | -34.7364 | -19.7648
11.8606 | -34.7369 | -19.7738
11.8667 | -34.7374 | -19.7828
11.8728 | -34.7378 | -19.7917
11.8789 | -34.7382 | -19.801
11.885 | -34.7386 | -19.8098
11.8911 | -34.7391 | -19.8189
11.8972 | -34.7395 | -19.8275
11.9033 | -34.7399 | -19.836
11.9094 | -34.7403 | -19.8445
11.9155 | -34.7407 | -19.8534
11.9216 | -34.7413 | -19.8622
11.9277 | -34.7417 | -19.8708
11.9338 | -34.7421 | -19.8798
11.9399 | -34.7426 | -19.889
11.946 | -34.743 | -19.8979
11.9521 | -34.7435 | -19.9067
11.9582 | -34.7439 | -19.9154
11.9644 | -34.7444 | -19.9241
11.9705 | -34.7449 | -19.9332
11.9766 | -34.7453 | -19.9421
11.9827 | -34.7457 | -19.9508
11.9888 | -34.7462 | -19.9596
11.9949 | -34.7466 | -19.9684
12.001 | -34.7471 | -19.9773
12.0071 | -34.7476 | -19.9858
12.0132 | -34.748 | -19.9944
12.0193 | -34.7484 | -20.0035
12.0254 | -34.7489 | -20.0121
12.0315 | -34.7493 | -20.0209
12.0376 | -34.7498 | -20.0295
12.0437 | -34.7502 | -20.0381
12.0498 | -34.7506 | -20.0471
12.0559 | -34.7511 | -20.056
12.062 | -34.7515 | -20.0648
12.0681 | -34.752 | -20.0736
12.0742 | -34.7525 | -20.0825
12.0803 | -34.7529 | -20.0913
12.0864 | -34.7534 | -20.1001
12.0925 | -34.7538 | -20.1089
12.0986 | -34.7543 | -20.1174
12.1047 | -34.7548 | -20.1262
12.1109 | -34.7553 | -20.1351
12.117 | -34.7557 | -20.1438
12.1231 | -34.7562 | -20.1526
12.1292 | -34.7566 | -20.1615
12.1353 | -34.7571 | -20.1701
12.1414 | -34.7576 | -20.1791
12.1475 | -34.758 | -20.1879
12.1536 | -34.7584 | -20.1967
12.1597 | -34.7589 | -20.2057
12.1658 | -34.7593 | -20.2145
12.1719 | -34.7598 | -20.2234
12.178 | -34.7603 | -20.2323
12.1841 | -34.7608 | -20.2412
12.1902 | -34.7612 | -20.25
12.1963 | -34.7617 | -20.2588
12.2024 | -34.7622 | -20.2675
12.2085 | -34.7626 | -20.2763
12.2146 | -34.763 | -20.2849
12.2207 | -34.7635 | -20.2937
12.2268 | -34.7639 | -20.3025
12.2329 | -34.7644 | -20.3112
12.239 | -34.7648 | -20.3198
12.2451 | -34.7653 | -20.3283
12.2513 | -34.7658 | -20.3372
12.2574 | -34.7662 | -20.3462
12.2635 | -34.7667 | -20.3546
12.2696 | -34.7671 | -20.3631
12.2757 | -34.7675 | -20.3716
12.2818 | -34.768 | -20.3802
12.2879 | -34.7684 | -20.3888
12.294 | -34.7688 | -20.3973
12.3001 | -34.7693 | -20.4062
12.3062 | -34.7698 | -20.415
12.3123 | -34.7702 | -20.4238
12.3184 | -34.7707 | -20.4323
12.3245 | -34.7712 | -20.4412
12.3306 | -34.7716 | -20.4496
12.3367 | -34.7721 | -20.4582
12.3428 | -34.7725 | -20.4667
12.3489 | -34.773 | -20.4749
12.355 | -34.7734 | -20.4836
12.3611 | -34.7739 | -20.4923
12.3672 | -34.7743 | -20.5009
12.3733 | -34.7748 | -20.51
12.3794 | -34.7752 | -20.5182
12.3855 | -34.7756 | -20.5267
12.3916 | -34.7761 | -20.5356
12.3978 | -34.7765 | -20.5446
12.4039 | -34.7769 | -20.5532
12.41 | -34.7774 | -20.5621
12.4161 | -34.7779 | -20.5705
12.4222 | -34.7784 | -20.579
12.4283 | -34.7789 | -20.5878
12.4344 | -34.7793 | -20.5964
12.4405 | -34.7797 | -20.6052
12.4466 | -34.7802 | -20.6137
12.4527 | -34.7806 | -20.6224
12.4588 | -34.7811 | -20.631
12.4649 | -34.7816 | -20.6399
12.471 | -34.782 | -20.6491
12.4771 | -34.7825 | -20.6571
12.4832 | -34.7829 | -20.6657
12.4893 | -34.7834 | -20.6741
12.4954 | -34.7839 | -20.6826
12.5015 | -34.7844 | -20.691
12.5076 | -34.7848 | -20.6997
12.5137 | -34.7852 | -20.7088
12.5198 | -34.7857 | -20.7171
12.5259 | -34.7862 | -20.7256
12.532 | -34.7866 | -20.7345
12.5382 | -34.7871 | -20.743
12.5443 | -34.7875 | -20.7517
12.5504 | -34.788 | -20.7604
12.5565 | -34.7885 | -20.7687
12.5626 | -34.789 | -20.7775
12.5687 | -34.7894 | -20.7864
12.5748 | -34.7898 | -20.7949
12.5809 | -34.7903 | -20.8031
12.587 | -34.7907 | -20.8121
12.5931 | -34.7912 | -20.8206
12.5992 | -34.7916 | -20.8295
12.6053 | -34.7921 | -20.8379
12.6114 | -34.7926 | -20.8465
12.6175 | -34.793 | -20.8553
12.6236 | -34.7935 | -20.8636
12.6297 | -34.794 | -20.8721
12.6358 | -34.7945 | -20.8804
12.6419 | -34.795 | -20.8892
12.648 | -34.7954 | -20.8981
12.6541 | -34.7959 | -20.9066
12.6602 | -34.7964 | -20.9147
12.6663 | -34.7968 | -20.9232
12.6724 | -34.7973 | -20.9318
12.6785 | -34.7977 | -20.9406
12.6847 | -34.7982 | -20.9492
12.6908 | -34.7986 | -20.9578
12.6969 | -34.7991 | -20.9663
12.703 | -34.7996 | -20.9749
12.7091 | -34.8 | -20.9836
12.7152 | -34.8005 | -20.9921
12.7213 | -34.8009 | -21.0008
12.7274 | -34.8014 | -21.0094
12.7335 | -34.8019 | -21.0173
12.7396 | -34.8023 | -21.0255
12.7457 | -34.8028 | -21.0341
12.7518 | -34.8032 | -21.0427
12.7579 | -34.8036 | -21.0515
12.764 | -34.8041 | -21.0601
12.7701 | -34.8044 | -21.0683
12.7762 | -34.8049 | -21.0769
12.7823 | -34.8054 | -21.0855
12.7884 | -34.8058 | -21.094
12.7945 | -34.8062 | -21.1022
12.8006 | -34.8067 | -21.1101
12.8067 | -34.8071 | -21.1188
12.8128 | -34.8076 | -21.1269
12.8189 | -34.808 | -21.1351
12.8251 | -34.8084 | -21.1434
12.8312 | -34.8088 | -21.1518
12.8373 | -34.8092 | -21.1603
12.8434 | -34.8098 | -21.1687
12.8495 | -34.8102 | -21.1771
12.8556 | -34.8106 | -21.1855
12.8617 | -34.811 | -21.1938
12.8678 | -34.8115 | -21.2022
12.8739 | -34.8119 | -21.2102
12.88 | -34.8124 | -21.2187
12.8861 | -34.8128 | -21.2272
12.8922 | -34.8132 | -21.2356
12.8983 | -34.8137 | -21.2437
12.9044 | -34.8142 | -21.2521
12.9105 | -34.8147 | -21.2604
12.9166 | -34.8152 | -21.2691
12.9227 | -34.8156 | -21.2781
12.9288 | -34.8161 | -21.2866
12.9349 | -34.8166 | -21.2954
12.941 | -34.8171 | -21.304
12.9471 | -34.8176 | -21.3123
12.9532 | -34.8181 | -21.3207
12.9593 | -34.8185 | -21.3293
12.9654 | -34.819 | -21.3379
12.9716 | -34.8194 | -21.3464
12.9777 | -34.8199 | -21.3544
12.9838 | -34.8203 | -21.3627
12.9899 | -34.8207 | -21.3712
12.996 | -34.8212 | -21.3799
13.0021 | -34.8216 | -21.3884
13.0082 | -34.822 | -21.3968
13.0143 | -34.8225 | -21.4048
13.0204 | -34.8229 | -21.4129
13.0265 | -34.8234 | -21.421
13.0326 | -34.8238 | -21.4296
13.0387 | -34.8243 | -21.4383
13.0448 | -34.8247 | -21.4466
13.0509 | -34.8252 | -21.455
13.057 | -34.8257 | -21.463
13.0631 | -34.8262 | -21.4712
13.0692 | -34.8267 | -21.4794
13.0753 | -34.8272 | -21.4878
13.0814 | -34.8276 | -21.4959
13.0875 | -34.828 | -21.5038
13.0936 | -34.8285 | -21.512
13.0997 | -34.829 | -21.5199
13.1058 | -34.8294 | -21.5277
13.112 | -34.8298 | -21.5361
13.1181 | -34.8303 | -21.5444
13.1242 | -34.8307 | -21.5528
13.1303 | -34.8312 | -21.5613
13.1364 | -34.8316 | -21.5697
13.1425 | -34.8321 | -21.5779
13.1486 | -34.8325 | -21.586
13.1547 | -34.833 | -21.594
13.1608 | -34.8334 | -21.6022
13.1669 | -34.8339 | -21.6105
13.173 | -34.8343 | -21.6187
13.1791 | -34.8348 | -21.6271
13.1852 | -34.8352 | -21.6356
13.1913 | -34.8357 | -21.6439
13.1974 | -34.8361 | -21.6521
13.2035 | -34.8365 | -21.6604
13.2096 | -34.837 | -21.6686
13.2157 | -34.8375 | -21.6769
13.2218 | -34.838 | -21.6852
13.2279 | -34.8385 | -21.6935
13.234 | -34.8389 | -21.7018
13.2401 | -34.8394 | -21.7099
13.2462 | -34.8398 | -21.7181
13.2524 | -34.8403 | -21.7262
13.2585 | -34.8407 | -21.7344
13.2646 | -34.8412 | -21.743
13.2707 | -34.8416 | -21.7513
13.2768 | -34.8421 | -21.7599
13.2829 | -34.8425 | -21.7683
13.289 | -34.843 | -21.7766
13.2951 | -34.8434 | -21.7848
13.3012 | -34.8438 | -21.7932
13.3073 | -34.8443 | -21.8019
13.3134 | -34.8447 | -21.8102
13.3195 | -34.8452 | -21.8185
13.3256 | -34.8456 | -21.827
13.3317 | -34.8461 | -21.8351
13.3378 | -34.8465 | -21.8435
13.3439 | -34.847 | -21.8518
13.35 | -34.8474 | -21.86
13.3561 | -34.8479 | -21.8683
13.3622 | -34.8483 | -21.8765
13.3683 | -34.8488 | -21.8845
13.3744 | -34.8492 | -21.8927
13.3805 | -34.8496 | -21.9011
13.3866 | -34.8501 | -21.909
13.3927 | -34.8505 | -21.9172
13.3989 | -34.851 | -21.9253
13.405 | -34.8514 | -21.9335
13.4111 | -34.8519 | -21.9417
13.4172 | -34.8524 | -21.9496
13.4233 | -34.8528 | -21.9581
13.4294 | -34.8532 | -21.966
13.4355 | -34.8537 | -21.9747
13.4416 | -34.8541 | -21.983
13.4477 | -34.8545 | -21.9913
13.4538 | -34.855 | -21.9994
13.4599 | -34.8555 | -22.0078
13.466 | -34.8559 | -22.0163
13.4721 | -34.8564 | -22.0248
13.4782 | -34.8568 | -22.033
13.4843 | -34.8573 | -22.0411
13.4904 | -34.8577 | -22.0488
13.4965 | -34.8581 | -22.0574
13.5026 | -34.8586 | -22.0654
13.5087 | -34.8591 | -22.0732
13.5148 | -34.8595 | -22.0816
13.5209 | -34.8599 | -22.0896
13.527 | -34.8604 | -22.098
13.5331 | -34.8608 | -22.1061
13.5393 | -34.8613 | -22.1144
13.5454 | -34.8617 | -22.1226
13.5515 | -34.8622 | -22.1308
13.5576 | -34.8626 | -22.1392
13.5637 | -34.8631 | -22.1473
13.5698 | -34.8635 | -22.1557
13.5759 | -34.8639 | -22.1635
13.582 | -34.8643 | -22.1718
13.5881 | -34.8648 | -22.18
13.5942 | -34.8652 | -22.1883
13.6003 | -34.8657 | -22.1966
13.6064 | -34.8661 | -22.205
13.6125 | -34.8667 | -22.2136
13.6186 | -34.8671 | -22.2218
13.6247 | -34.8676 | -22.23
13.6308 | -34.868 | -22.2379
13.6369 | -34.8685 | -22.2461
13.643 | -34.8689 | -22.2542
13.6491 | -34.8694 | -22.2626
13.6552 | -34.8698 | -22.2706
13.6613 | -34.8703 | -22.279
13.6674 | -34.8707 | -22.2873
13.6735 | -34.8712 | -22.2956
13.6796 | -34.8717 | -22.3037
13.6858 | -34.872 | -22.3119
13.6919 | -34.8725 | -22.3202
13.698 | -34.8729 | -22.3285
13.7041 | -34.8733 | -22.337
13.7102 | -34.8739 | -22.3453
13.7163 | -34.8744 | -22.3535
13.7224 | -34.8748 | -22.3616
13.7285 | -34.8753 | -22.3699
13.7346 | -34.8757 | -22.378
13.7407 | -34.8762 | -22.3863
13.7468 | -34.8766 | -22.3945
13.7529 | -34.877 | -22.4026
13.759 | -34.8775 | -22.411
13.7651 | -34.8779 | -22.4193
13.7712 | -34.8783 | -22.4277
13.7773 | -34.8788 | -22.436
13.7834 | -34.8793 | -22.4441
13.7895 | -34.8798 | -22.4523
13.7956 | -34.8802 | -22.4606
13.8017 | -34.8807 | -22.4687
13.8078 | -34.8812 | -22.477
13.8139 | -34.8816 | -22.485
13.82 | -34.8821 | -22.4929
13.8262 | -34.8826 | -22.5016
13.8323 | -34.883 | -22.5099
13.8384 | -34.8835 | -22.5183
13.8445 | -34.884 | -22.5265
13.8506 | -34.8844 | -22.5349
13.8567 | -34.8848 | -22.5434
13.8628 | -34.8853 | -22.5518
13.8689 | -34.8857 | -22.5601
13.875 | -34.8862 | -22.5682
13.8811 | -34.8867 | -22.5763
13.8872 | -34.8871 | -22.5845
13.8933 | -34.8876 | -22.5926
13.8994 | -34.8881 | -22.6006
13.9055 | -34.8885 | -22.6087
13.9116 | -34.889 | -22.6167
13.9177 | -34.8895 | -22.6246
13.9238 | -34.89 | -22.633
13.9299 | -34.8905 | -22.6412
13.936 | -34.8909 | -22.6493
13.9421 | -34.8913 | -22.6575
13.9482 | -34.8918 | -22.6657
13.9543 | -34.8923 | -22.6737
13.9604 | -34.8927 | -22.682
13.9665 | -34.8932 | -22.6903
13.9727 | -34.8937 | -22.6983
13.9788 | -34.8941 | -22.7068
13.9849 | -34.8945 | -22.7151
13.991 | -34.8949 | -22.7235
13.9971 | -34.8954 | -22.732
14.0032 | -34.8958 | -22.7404
14.0093 | -34.8962 | -22.7485
14.0154 | -34.8966 | -22.757
14.0215 | -34.8971 | -22.7649
14.0276 | -34.8976 | -22.773
14.0337 | -34.8981 | -22.7815
14.0398 | -34.8985 | -22.7895
14.0459 | -34.8989 | -22.7979
14.052 | -34.8993 | -22.8061
14.0581 | -34.8998 | -22.814
14.0642 | -34.9003 | -22.8224
14.0703 | -34.9007 | -22.8307
14.0764 | -34.9012 | -22.8393
14.0825 | -34.9017 | -22.8476
14.0886 | -34.9021 | -22.8554
14.0947 | -34.9025 | -22.8634
14.1008 | -34.903 | -22.8719
14.1069 | -34.9034 | -22.8801
14.1131 | -34.9039 | -22.8886
14.1192 | -34.9043 | -22.8965
14.1253 | -34.9047 | -22.9044
14.1314 | -34.9052 | -22.9126
14.1375 | -34.9057 | -22.921
14.1436 | -34.9062 | -22.9292
14.1497 | -34.9067 | -22.9371
14.1558 | -34.9071 | -22.9454
14.1619 | -34.9076 | -22.9536
14.168 | -34.908 | -22.9617
14.1741 | -34.9086 | -22.9698
14.1802 | -34.909 | -22.9778
14.1863 | -34.9094 | -22.9861
14.1924 | -34.9099 | -22.9943
14.1985 | -34.9104 | -23.0022
14.2046 | -34.9108 | -23.0102
14.2107 | -34.9113 | -23.0181
14.2168 | -34.9118 | -23.0267
14.2229 | -34.9122 | -23.0349
14.229 | -34.9127 | -23.0431
14.2351 | -34.9132 | -23.051
14.2412 | -34.9136 | -23.059
14.2473 | -34.9141 | -23.0673
14.2534 | -34.9146 | -23.0754
14.2596 | -34.915 | -23.0836
14.2657 | -34.9155 | -23.0919
14.2718 | -34.916 | -23.1003
14.2779 | -34.9164 | -23.1083
14.284 | -34.9169 | -23.1165
14.2901 | -34.9174 | -23.1244
14.2962 | -34.9179 | -23.1323
14.3023 | -34.9183 | -23.1406
14.3084 | -34.9188 | -23.1491
14.3145 | -34.9193 | -23.1572
14.3206 | -34.9197 | -23.1655
14.3267 | -34.9202 | -23.1736
14.3328 | -34.9207 | -23.1817
14.3389 | -34.9211 | -23.1896
14.345 | -34.9215 | -23.1979
14.3511 | -34.922 | -23.2059
14.3572 | -34.9224 | -23.2143
14.3633 | -34.9229 | -23.2223
14.3694 | -34.9233 | -23.2304
14.3755 | -34.9238 | -23.2383
14.3816 | -34.9243 | -23.2463
14.3877 | -34.9248 | -23.2543
14.3938 | -34.9252 | -23.2627
14.4 | -34.9257 | -23.2707
14.4061 | -34.9261 | -23.2788
14.4122 | -34.9265 | -23.2872
14.4183 | -34.927 | -23.2953
14.4244 | -34.9274 | -23.3033
14.4305 | -34.9279 | -23.3118
14.4366 | -34.9284 | -23.32
14.4427 | -34.9289 | -23.3284
14.4488 | -34.9293 | -23.3362
14.4549 | -34.9298 | -23.3443
14.461 | -34.9302 | -23.3526
14.4671 | -34.9307 | -23.3607
14.4732 | -34.9312 | -23.3685
14.4793 | -34.9316 | -23.3773
14.4854 | -34.9322 | -23.385
14.4915 | -34.9326 | -23.393
14.4976 | -34.9331 | -23.4012
14.5037 | -34.9336 | -23.4095
14.5098 | -34.9341 | -23.4174
14.5159 | -34.9346 | -23.4258
14.522 | -34.935 | -23.434
14.5281 | -34.9355 | -23.4419
14.5342 | -34.936 | -23.4502
14.5403 | -34.9364 | -23.4584
14.5465 | -34.9369 | -23.4669
14.5526 | -34.9374 | -23.4752
14.5587 | -34.9378 | -23.4834
14.5648 | -34.9383 | -23.4913
14.5709 | -34.9388 | -23.4996
14.577 | -34.9392 | -23.5078
14.5831 | -34.9397 | -23.5161
14.5892 | -34.9403 | -23.5238
14.5953 | -34.9407 | -23.5316
14.6014 | -34.9412 | -23.5397
14.6075 | -34.9417 | -23.5476
14.6136 | -34.9421 | -23.5555
14.6197 | -34.9426 | -23.5637
14.6258 | -34.943 | -23.5719
14.6319 | -34.9435 | -23.5799
14.638 | -34.944 | -23.588
14.6441 | -34.9445 | -23.5956
14.6502 | -34.9449 | -23.6035
14.6563 | -34.9454 | -23.6116
14.6624 | -34.9459 | -23.6196
14.6685 | -34.9464 | -23.6273
14.6746 | -34.9468 | -23.6354
14.6807 | -34.9473 | -23.6435
14.6869 | -34.9478 | -23.6515
14.693 | -34.9483 | -23.6595
14.6991 | -34.9487 | -23.6674
14.7052 | -34.9492 | -23.6753
14.7113 | -34.9497 | -23.6836
14.7174 | -34.9501 | -23.6916
14.7235 | -34.9506 | -23.6995
14.7296 | -34.9511 | -23.7075
14.7357 | -34.9516 | -23.7154
14.7418 | -34.952 | -23.7232
14.7479 | -34.9525 | -23.7313
14.754 | -34.953 | -23.7393
14.7601 | -34.9535 | -23.7474
14.7662 | -34.954 | -23.7555
14.7723 | -34.9544 | -23.7634
14.7784 | -34.9549 | -23.7716
14.7845 | -34.9554 | -23.7797
14.7906 | -34.9559 | -23.7878
14.7967 | -34.9563 | -23.7957
14.8028 | -34.9568 | -23.8036
14.8089 | -34.9573 | -23.8115
14.815 | -34.9577 | -23.8193
14.8211 | -34.9581 | -23.8272
14.8272 | -34.9586 | -23.8354
14.8334 | -34.9591 | -23.8433
14.8395 | -34.9596 | -23.8514
14.8456 | -34.9601 | -23.8592
14.8517 | -34.9606 | -23.8673
14.8578 | -34.961 | -23.8752
14.8639 | -34.9615 | -23.8834
14.87 | -34.962 | -23.8915
14.8761 | -34.9625 | -23.8992
14.8822 | -34.963 | -23.9073
14.8883 | -34.9634 | -23.9151
14.8944 | -34.9639 | -23.9232
14.9005 | -34.9644 | -23.931
14.9066 | -34.9649 | -23.939
14.9127 | -34.9654 | -23.9468
14.9188 | -34.9659 | -23.955
14.9249 | -34.9664 | -23.963
14.931 | -34.9668 | -23.9708
14.9371 | -34.9673 | -23.9786
14.9432 | -34.9678 | -23.9868
14.9493 | -34.9683 | -23.9947
14.9554 | -34.9688 | -24.0026
14.9615 | -34.9692 | -24.0107
14.9676 | -34.9697 | -24.0189
14.9738 | -34.9702 | -24.0267
14.9799 | -34.9707 | -24.0351
14.986 | -34.9712 | -24.0427
14.9921 | -34.9717 | -24.0505
14.9982 | -34.9722 | -24.0588
15.0043 | -34.9726 | -24.0667
15.0104 | -34.9731 | -24.0746
15.0165 | -34.9736 | -24.0827
15.0226 | -34.9741 | -24.0904
15.0287 | -34.9745 | -24.0988
15.0348 | -34.9751 | -24.1068
15.0409 | -34.9755 | -24.1151
15.047 | -34.976 | -24.1233
15.0531 | -34.9766 | -24.1314
15.0592 | -34.9771 | -24.1395
15.0653 | -34.9776 | -24.1474
15.0714 | -34.9781 | -24.1552
15.0775 | -34.9786 | -24.1629
15.0836 | -34.9791 | -24.1708
15.0897 | -34.9796 | -24.1787
15.0958 | -34.9801 | -24.1865
15.1019 | -34.9805 | -24.1944
15.108 | -34.981 | -24.2024
15.1141 | -34.9815 | -24.2103
15.1203 | -34.982 | -24.2183
15.1264 | -34.9826 | -24.2265
15.1325 | -34.9831 | -24.2343
15.1386 | -34.9836 | -24.2424
15.1447 | -34.9841 | -24.2503
15.1508 | -34.9846 | -24.2586
15.1569 | -34.9851 | -24.2668
15.163 | -34.9857 | -24.2744
15.1691 | -34.9862 | -24.2823
15.1752 | -34.9867 | -24.2902
15.1813 | -34.9872 | -24.2981
15.1874 | -34.9877 | -24.3062
15.1935 | -34.9883 | -24.3139
15.1996 | -34.9888 | -24.3221
15.2057 | -34.9893 | -24.3298
15.2118 | -34.9897 | -24.3377
15.2179 | -34.9903 | -24.3454
15.224 | -34.9907 | -24.3535
15.2301 | -34.9912 | -24.3611
15.2362 | -34.9918 | -24.3691
15.2423 | -34.9923 | -24.3773
15.2484 | -34.9927 | -24.385
15.2545 | -34.9932 | -24.3929
15.2607 | -34.9938 | -24.4008
15.2668 | -34.9943 | -24.4084
15.2729 | -34.9949 | -24.4162
15.279 | -34.9954 | -24.4242
15.2851 | -34.9959 | -24.432
15.2912 | -34.9964 | -24.4397
15.2973 | -34.9968 | -24.4476
15.3034 | -34.9974 | -24.4552
15.3095 | -34.9979 | -24.4628
15.3156 | -34.9984 | -24.4709
15.3217 | -34.9989 | -24.4783
15.3278 | -34.9994 | -24.4863
15.3339 | -34.9999 | -24.4939
15.34 | -35.0005 | -24.5016
15.3461 | -35.0009 | -24.5094
15.3522 | -35.0014 | -24.5172
15.3583 | -35.002 | -24.525
15.3644 | -35.0025 | -24.533
15.3705 | -35.003 | -24.5406
15.3766 | -35.0035 | -24.5488
15.3827 | -35.004 | -24.5563
15.3888 | -35.0045 | -24.5639
15.3949 | -35.005 | -24.5716
15.401 | -35.0055 | -24.5795
15.4072 | -35.006 | -24.587
15.4133 | -35.0065 | -24.5948
15.4194 | -35.0071 | -24.6027
15.4255 | -35.0076 | -24.6105
15.4316 | -35.0081 | -24.6184
15.4377 | -35.0086 | -24.6261
15.4438 | -35.0092 | -24.6336
15.4499 | -35.0096 | -24.6414
15.456 | -35.0101 | -24.6492
15.4621 | -35.0106 | -24.657
15.4682 | -35.0112 | -24.6649
15.4743 | -35.0117 | -24.6725
15.4804 | -35.0122 | -24.6805
15.4865 | -35.0127 | -24.6883
15.4926 | -35.0132 | -24.6962
15.4987 | -35.0138 | -24.7041
15.5048 | -35.0143 | -24.7118
15.5109 | -35.0148 | -24.7192
15.517 | -35.0153 | -24.7272
15.5231 | -35.0159 | -24.7349
15.5292 | -35.0164 | -24.7425
15.5353 | -35.0169 | -24.7503
15.5414 | -35.0174 | -24.7584
15.5476 | -35.0179 | -24.7663
15.5537 | -35.0185 | -24.7739
15.5598 | -35.019 | -24.7815
15.5659 | -35.0195 | -24.7895
15.572 | -35.02 | -24.7968
15.5781 | -35.0205 | -24.8046
15.5842 | -35.0211 | -24.8121
15.5903 | -35.0216 | -24.8195
15.5964 | -35.0222 | -24.8272
15.6025 | -35.0227 | -24.8348
15.6086 | -35.0232 | -24.8426
15.6147 | -35.0237 | -24.8503
15.6208 | -35.0242 | -24.8578
15.6269 | -35.0248 | -24.8654
15.633 | -35.0253 | -24.8732
15.6391 | -35.0258 | -24.8811
15.6452 | -35.0263 | -24.8887
15.6513 | -35.0269 | -24.8965
15.6574 | -35.0274 | -24.904
15.6635 | -35.0279 | -24.9117
15.6696 | -35.0284 | -24.9193
15.6757 | -35.029 | -24.9274
15.6818 | -35.0295 | -24.9353
15.688 | -35.0301 | -24.943
15.6941 | -35.0306 | -24.9507
15.7002 | -35.0311 | -24.9584
15.7063 | -35.0316 | -24.9662
15.7124 | -35.0321 | -24.974
15.7185 | -35.0327 | -24.9817
15.7246 | -35.0332 | -24.9892
15.7307 | -35.0337 | -24.997
15.7368 | -35.0343 | -25.0051
15.7429 | -35.0348 | -25.0127
15.749 | -35.0353 | -25.0204
15.7551 | -35.0358 | -25.028
15.7612 | -35.0363 | -25.0352
15.7673 | -35.0368 | -25.0428
15.7734 | -35.0373 | -25.0503
15.7795 | -35.0378 | -25.0585
15.7856 | -35.0384 | -25.0659
15.7917 | -35.0389 | -25.0734
15.7978 | -35.0394 | -25.081
15.8039 | -35.04 | -25.0887
15.81 | -35.0405 | -25.0961
15.8161 | -35.041 | -25.1039
15.8222 | -35.0415 | -25.1116
15.8283 | -35.0421 | -25.119
15.8345 | -35.0426 | -25.1267
15.8406 | -35.0432 | -25.1344
15.8467 | -35.0437 | -25.1421
15.8528 | -35.0442 | -25.1497
15.8589 | -35.0448 | -25.1573
15.865 | -35.0453 | -25.1648
15.8711 | -35.0458 | -25.1728
15.8772 | -35.0464 | -25.1805
15.8833 | -35.0469 | -25.1886
15.8894 | -35.0474 | -25.1962
15.8955 | -35.048 | -25.204
15.9016 | -35.0485 | -25.2115
15.9077 | -35.0491 | -25.2192
15.9138 | -35.0496 | -25.2267
15.9199 | -35.0501 | -25.2345
15.926 | -35.0507 | -25.2425
15.9321 | -35.0512 | -25.2501
15.9382 | -35.0518 | -25.2575
15.9443 | -35.0524 | -25.2654
15.9504 | -35.0529 | -25.273
15.9565 | -35.0535 | -25.2806
15.9626 | -35.054 | -25.2884
15.9687 | -35.0546 | -25.2958
15.9749 | -35.0551 | -25.3034
15.981 | -35.0557 | -25.311
15.9871 | -35.0562 | -25.3187
15.9932 | -35.0567 | -25.3265
15.9993 | -35.0573 | -25.334
16.0054 | -35.0578 | -25.3413
16.0115 | -35.0584 | -25.3488
16.0176 | -35.0589 | -25.3567
16.0237 | -35.0594 | -25.3647
16.0298 | -35.06 | -25.3718
16.0359 | -35.0605 | -25.3793
16.042 | -35.0611 | -25.3868
16.0481 | -35.0616 | -25.3945
16.0542 | -35.0621 | -25.4024
16.0603 | -35.0627 | -25.4099
16.0664 | -35.0632 | -25.4175
16.0725 | -35.0638 | -25.425
16.0786 | -35.0644 | -25.4328
16.0847 | -35.0648 | -25.4401
16.0908 | -35.0654 | -25.4478
16.0969 | -35.0659 | -25.4552
16.103 | -35.0664 | -25.4629
16.1091 | -35.067 | -25.4704
16.1152 | -35.0675 | -25.4781
16.1214 | -35.0681 | -25.4858
16.1275 | -35.0686 | -25.4934
16.1336 | -35.0691 | -25.501
16.1397 | -35.0697 | -25.5084
16.1458 | -35.0702 | -25.5159
16.1519 | -35.0708 | -25.5237
16.158 | -35.0713 | -25.5313
16.1641 | -35.0718 | -25.5389
16.1702 | -35.0724 | -25.5467
16.1763 | -35.0729 | -25.5543
16.1824 | -35.0735 | -25.5616
16.1885 | -35.074 | -25.5692
16.1946 | -35.0746 | -25.5767
16.2007 | -35.0751 | -25.5846
16.2068 | -35.0757 | -25.592
16.2129 | -35.0762 | -25.5997
16.219 | -35.0767 | -25.6071
16.2251 | -35.0772 | -25.615
16.2312 | -35.0777 | -25.6226
16.2373 | -35.0783 | -25.6302
16.2434 | -35.0788 | -25.6377
16.2495 | -35.0794 | -25.645
16.2556 | -35.08 | -25.6527
16.2618 | -35.0805 | -25.6601
16.2679 | -35.0811 | -25.6676
16.274 | -35.0816 | -25.6753
16.2801 | -35.0822 | -25.6826
16.2862 | -35.0827 | -25.69
16.2923 | -35.0832 | -25.6977
16.2984 | -35.0838 | -25.7054
16.3045 | -35.0844 | -25.7129
16.3106 | -35.0849 | -25.7201
16.3167 | -35.0855 | -25.7278
16.3228 | -35.086 | -25.7355
16.3289 | -35.0865 | -25.7431
16.335 | -35.0871 | -25.7509
16.3411 | -35.0876 | -25.7582
16.3472 | -35.0881 | -25.7661
16.3533 | -35.0887 | -25.7737
16.3594 | -35.0892 | -25.7811
16.3655 | -35.0897 | -25.7888
16.3716 | -35.0902 | -25.7963
16.3777 | -35.0908 | -25.8038
16.3838 | -35.0913 | -25.8113
16.3899 | -35.0918 | -25.8189
16.396 | -35.0923 | -25.8267
16.4021 | -35.0929 | -25.8343
16.4083 | -35.0935 | -25.8417
16.4144 | -35.094 | -25.8491
16.4205 | -35.0946 | -25.8568
16.4266 | -35.0951 | -25.8643
16.4327 | -35.0957 | -25.8717
16.4388 | -35.0962 | -25.8792
16.4449 | -35.0967 | -25.8868
16.451 | -35.0973 | -25.8945
16.4571 | -35.0978 | -25.9023
16.4632 | -35.0983 | -25.9095
16.4693 | -35.0988 | -25.9171
16.4754 | -35.0994 | -25.9247
16.4815 | -35.0999 | -25.9324
16.4876 | -35.1005 | -25.9399
16.4937 | -35.101 | -25.9474
16.4998 | -35.1016 | -25.955
16.5059 | -35.1021 | -25.9624
16.512 | -35.1027 | -25.9702
16.5181 | -35.1032 | -25.9779
16.5242 | -35.1037 | -25.9854
16.5303 | -35.1043 | -25.9927
16.5364 | -35.1048 | -26.0001
16.5425 | -35.1053 | -26.0077
16.5487 | -35.1058 | -26.0152
16.5548 | -35.1064 | -26.0226
16.5609 | -35.1069 | -26.0304
16.567 | -35.1074 | -26.0379
16.5731 | -35.108 | -26.0455
16.5792 | -35.1085 | -26.053
16.5853 | -35.1091 | -26.0604
16.5914 | -35.1096 | -26.0681
16.5975 | -35.1101 | -26.0758
16.6036 | -35.1107 | -26.0832
16.6097 | -35.1112 | -26.0907
16.6158 | -35.1117 | -26.0984
16.6219 | -35.1123 | -26.1058
16.628 | -35.1128 | -26.1133
16.6341 | -35.1133 | -26.1206
16.6402 | -35.1139 | -26.1275
16.6463 | -35.1144 | -26.135
16.6524 | -35.1149 | -26.1424
16.6585 | -35.1155 | -26.1501
16.6646 | -35.116 | -26.1577
16.6707 | -35.1166 | -26.1648
16.6768 | -35.1171 | -26.1723
16.6829 | -35.1177 | -26.1798
16.689 | -35.1182 | -26.1874
16.6952 | -35.1187 | -26.1951
16.7013 | -35.1192 | -26.2026
16.7074 | -35.1198 | -26.2101
16.7135 | -35.1203 | -26.2175
16.7196 | -35.1208 | -26.2251
16.7257 | -35.1214 | -26.2324
16.7318 | -35.1219 | -26.2399
16.7379 | -35.1224 | -26.2475
16.744 | -35.123 | -26.2546
16.7501 | -35.1235 | -26.2622
16.7562 | -35.1241 | -26.2697
16.7623 | -35.1246 | -26.2772
16.7684 | -35.1251 | -26.2848
16.7745 | -35.1257 | -26.2921
16.7806 | -35.1262 | -26.2997
16.7867 | -35.1267 | -26.3071
16.7928 | -35.1272 | -26.3146
16.7989 | -35.1277 | -26.3221
16.805 | -35.1283 | -26.3296
16.8111 | -35.1288 | -26.3371
16.8172 | -35.1293 | -26.3449
16.8233 | -35.1299 | -26.3526
16.8294 | -35.1304 | -26.3602
16.8356 | -35.1309 | -26.3679
16.8417 | -35.1315 | -26.3753
16.8478 | -35.132 | -26.3826
16.8539 | -35.1325 | -26.39
16.86 | -35.133 | -26.3974
16.8661 | -35.1336 | -26.4049
16.8722 | -35.1341 | -26.4124
16.8783 | -35.1347 | -26.4196
16.8844 | -35.1352 | -26.427
16.8905 | -35.1358 | -26.4347
16.8966 | -35.1362 | -26.4422
16.9027 | -35.1368 | -26.4495
16.9088 | -35.1373 | -26.4573
16.9149 | -35.1379 | -26.4651
16.921 | -35.1384 | -26.4724
16.9271 | -35.139 | -26.4801
16.9332 | -35.1395 | -26.4877
16.9393 | -35.14 | -26.4955
16.9454 | -35.1405 | -26.5028
16.9515 | -35.141 | -26.5105
16.9576 | -35.1415 | -26.5179
16.9637 | -35.1421 | -26.5254
16.9698 | -35.1427 | -26.5325
16.9759 | -35.1432 | -26.5398
16.9821 | -35.1437 | -26.5475
16.9882 | -35.1442 | -26.5549
16.9943 | -35.1448 | -26.5625
17.0004 | -35.1453 | -26.57
17.0065 | -35.1458 | -26.5776
17.0126 | -35.1463 | -26.5849
17.0187 | -35.1469 | -26.5926
17.0248 | -35.1474 | -26.5997
17.0309 | -35.1479 | -26.6072
17.037 | -35.1484 | -26.6146
17.0431 | -35.149 | -26.6218
17.0492 | -35.1495 | -26.6292
17.0553 | -35.15 | -26.6369
17.0614 | -35.1505 | -26.6443
17.0675 | -35.151 | -26.652
17.0736 | -35.1516 | -26.6596
17.0797 | -35.1521 | -26.6667
17.0858 | -35.1526 | -26.6743
17.0919 | -35.1531 | -26.682
17.098 | -35.1537 | -26.6893
17.1041 | -35.1542 | -26.6966
17.1102 | -35.1547 | -26.7042
17.1163 | -35.1553 | -26.7117
17.1225 | -35.1558 | -26.7189
17.1286 | -35.1563 | -26.7261
17.1347 | -35.1569 | -26.7335
17.1408 | -35.1574 | -26.7409
17.1469 | -35.158 | -26.7482
17.153 | -35.1585 | -26.7556
17.1591 | -35.159 | -26.7628
17.1652 | -35.1595 | -26.77
17.1713 | -35.16 | -26.7774
17.1774 | -35.1606 | -26.7849
17.1835 | -35.1611 | -26.7925
17.1896 | -35.1616 | -26.7998
17.1957 | -35.1621 | -26.8074
17.2018 | -35.1626 | -26.8148
17.2079 | -35.1631 | -26.8222
17.214 | -35.1636 | -26.8295
17.2201 | -35.1641 | -26.8368
17.2262 | -35.1646 | -26.8444
17.2323 | -35.1652 | -26.8517
17.2384 | -35.1657 | -26.8591
17.2445 | -35.1662 | -26.8664
17.2506 | -35.1667 | -26.8739
17.2567 | -35.1672 | -26.8814
17.2628 | -35.1677 | -26.8891
17.269 | -35.1682 | -26.8965
17.2751 | -35.1687 | -26.9038
17.2812 | -35.1693 | -26.9111
17.2873 | -35.1698 | -26.9186
17.2934 | -35.1703 | -26.926
17.2995 | -35.1708 | -26.9332
17.3056 | -35.1713 | -26.9405
17.3117 | -35.1718 | -26.9478
17.3178 | -35.1723 | -26.9551
17.3239 | -35.1728 | -26.9625
17.33 | -35.1733 | -26.9698
17.3361 | -35.1738 | -26.9771
17.3422 | -35.1744 | -26.9845
17.3483 | -35.1749 | -26.9919
17.3544 | -35.1754 | -26.9992
17.3605 | -35.1759 | -27.0066
17.3666 | -35.1765 | -27.0142
17.3727 | -35.177 | -27.0215
17.3788 | -35.1775 | -27.0286
17.3849 | -35.178 | -27.0359
17.391 | -35.1785 | -27.0431
17.3971 | -35.1791 | -27.0505
17.4032 | -35.1796 | -27.0582
17.4094 | -35.1801 | -27.0655
17.4155 | -35.1805 | -27.0727
17.4216 | -35.181 | -27.08
17.4277 | -35.1815 | -27.0876
17.4338 | -35.1821 | -27.095
17.4399 | -35.1826 | -27.1025
17.446 | -35.1831 | -27.1096
17.4521 | -35.1836 | -27.1168
17.4582 | -35.1841 | -27.1241
17.4643 | -35.1846 | -27.1315
17.4704 | -35.1851 | -27.1387
17.4765 | -35.1857 | -27.1462
17.4826 | -35.1862 | -27.1534
17.4887 | -35.1867 | -27.161
17.4948 | -35.1872 | -27.1682
17.5009 | -35.1877 | -27.1754
17.507 | -35.1882 | -27.1828
17.5131 | -35.1887 | -27.1901
17.5192 | -35.1892 | -27.1974
17.5253 | -35.1897 | -27.2048
17.5314 | -35.1902 | -27.2121
17.5375 | -35.1907 | -27.2196
17.5436 | -35.1912 | -27.2271
17.5497 | -35.1917 | -27.2345
17.5559 | -35.1922 | -27.2417
17.562 | -35.1927 | -27.2489
17.5681 | -35.1932 | -27.2562
17.5742 | -35.1937 | -27.2635
17.5803 | -35.1943 | -27.2707
17.5864 | -35.1948 | -27.2778
17.5925 | -35.1953 | -27.2854
17.5986 | -35.1959 | -27.2925
17.6047 | -35.1964 | -27.3002
17.6108 | -35.1969 | -27.3073
17.6169 | -35.1974 | -27.3145
17.623 | -35.1979 | -27.3217
17.6291 | -35.1984 | -27.329
17.6352 | -35.1989 | -27.3361
17.6413 | -35.1994 | -27.3436
17.6474 | -35.1999 | -27.351
17.6535 | -35.2004 | -27.3584
17.6596 | -35.2009 | -27.3657
17.6657 | -35.2014 | -27.3729
17.6718 | -35.2019 | -27.3804
17.6779 | -35.2024 | -27.3876
17.684 | -35.2029 | -27.395
17.6901 | -35.2035 | -27.4022
17.6963 | -35.204 | -27.4093
17.7024 | -35.2045 | -27.4164
17.7085 | -35.205 | -27.4235
17.7146 | -35.2055 | -27.4308
17.7207 | -35.206 | -27.438
17.7268 | -35.2065 | -27.4452
17.7329 | -35.207 | -27.4527
17.739 | -35.2075 | -27.4602
17.7451 | -35.2081 | -27.4674
17.7512 | -35.2086 | -27.4748
17.7573 | -35.209 | -27.4822
17.7634 | -35.2095 | -27.4892
17.7695 | -35.21 | -27.4966
17.7756 | -35.2105 | -27.5039
17.7817 | -35.211 | -27.511
17.7878 | -35.2115 | -27.5183
17.7939 | -35.2121 | -27.5254
17.8 | -35.2126 | -27.5325
17.8061 | -35.2131 | -27.5399
17.8122 | -35.2136 | -27.5471
17.8183 | -35.2141 | -27.5545
17.8244 | -35.2146 | -27.5617
17.8305 | -35.2151 | -27.569
17.8366 | -35.2156 | -27.576
17.8428 | -35.2161 | -27.5836
17.8489 | -35.2166 | -27.5907
17.855 | -35.2171 | -27.5978
17.8611 | -35.2177 | -27.6048
17.8672 | -35.2182 | -27.6119
17.8733 | -35.2187 | -27.6194
17.8794 | -35.2192 | -27.6269
17.8855 | -35.2197 | -27.634
17.8916 | -35.2202 | -27.641
17.8977 | -35.2207 | -27.6483
17.9038 | -35.2212 | -27.6553
17.9099 | -35.2217 | -27.6625
17.916 | -35.2222 | -27.6697
17.9221 | -35.2227 | -27.677
17.9282 | -35.2232 | -27.6842
17.9343 | -35.2237 | -27.6915
17.9404 | -35.2242 | -27.6986
17.9465 | -35.2247 | -27.706
17.9526 | -35.2252 | -27.7134
17.9587 | -35.2257 | -27.7206
17.9648 | -35.2262 | -27.7278
17.9709 | -35.2267 | -27.735
17.977 | -35.2272 | -27.742
17.9832 | -35.2277 | -27.7493
17.9893 | -35.2281 | -27.7566
17.9954 | -35.2287 | -27.764
18.0015 | -35.2292 | -27.7713
18.0076 | -35.2297 | -27.778
18.0137 | -35.2302 | -27.7855
18.0198 | -35.2306 | -27.7926
18.0259 | -35.2311 | -27.8
18.032 | -35.2316 | -27.807
18.0381 | -35.2321 | -27.814
18.0442 | -35.2326 | -27.8213
18.0503 | -35.2331 | -27.8287
18.0564 | -35.2336 | -27.8359
18.0625 | -35.2341 | -27.8433
18.0686 | -35.2346 | -27.8504
18.0747 | -35.2351 | -27.8575
18.0808 | -35.2356 | -27.8645
18.0869 | -35.2361 | -27.8718
18.093 | -35.2366 | -27.879
18.0991 | -35.2371 | -27.8865
18.1052 | -35.2376 | -27.8939
18.1113 | -35.2381 | -27.9008
18.1174 | -35.2386 | -27.9078
18.1236 | -35.2392 | -27.915
18.1297 | -35.2397 | -27.9222
18.1358 | -35.2402 | -27.9294
18.1419 | -35.2408 | -27.9367
18.148 | -35.2413 | -27.9438
18.1541 | -35.2419 | -27.9507
18.1602 | -35.2424 | -27.958
18.1663 | -35.2429 | -27.9652
18.1724 | -35.2434 | -27.9726
18.1785 | -35.2439 | -27.9799
18.1846 | -35.2444 | -27.9869
18.1907 | -35.2449 | -27.9944
18.1968 | -35.2453 | -28.0017
18.2029 | -35.2458 | -28.0085
18.209 | -35.2464 | -28.016
18.2151 | -35.2469 | -28.0235
18.2212 | -35.2474 | -28.0304
18.2273 | -35.248 | -28.0375
18.2334 | -35.2485 | -28.0446
18.2395 | -35.2489 | -28.0516
18.2456 | -35.2494 | -28.0586
18.2517 | -35.2499 | -28.0658
18.2578 | -35.2504 | -28.0727
18.2639 | -35.2509 | -28.0802
18.2701 | -35.2515 | -28.0873
18.2762 | -35.2519 | -28.0946
18.2823 | -35.2525 | -28.1019
18.2884 | -35.253 | -28.1091
18.2945 | -35.2535 | -28.1161
18.3006 | -35.254 | -28.1234
18.3067 | -35.2545 | -28.1306
18.3128 | -35.255 | -28.1377
18.3189 | -35.2555 | -28.1451
18.325 | -35.256 | -28.1517
18.3311 | -35.2565 | -28.159
18.3372 | -35.257 | -28.1659
18.3433 | -35.2575 | -28.1728
18.3494 | -35.258 | -28.1799
18.3555 | -35.2585 | -28.1869
18.3616 | -35.259 | -28.1938
18.3677 | -35.2596 | -28.2007
18.3738 | -35.26 | -28.2077
18.3799 | -35.2605 | -28.2146
18.386 | -35.261 | -28.2216
18.3921 | -35.2615 | -28.2286
18.3982 | -35.262 | -28.2355
18.4043 | -35.2624 | -28.2427
18.4105 | -35.2629 | -28.2499
18.4166 | -35.2634 | -28.2569
18.4227 | -35.2639 | -28.2641
18.4288 | -35.2645 | -28.2714
18.4349 | -35.265 | -28.2784
18.441 | -35.2655 | -28.2855
18.4471 | -35.266 | -28.2927
18.4532 | -35.2665 | -28.2998
18.4593 | -35.267 | -28.3068
18.4654 | -35.2675 | -28.3136
18.4715 | -35.268 | -28.3207
18.4776 | -35.2685 | -28.3279
18.4837 | -35.269 | -28.3351
18.4898 | -35.2695 | -28.3424
18.4959 | -35.2701 | -28.3496
18.502 | -35.2706 | -28.3568
18.5081 | -35.2711 | -28.3642
18.5142 | -35.2715 | -28.3714
18.5203 | -35.2721 | -28.3783
18.5264 | -35.2726 | -28.385
18.5325 | -35.2731 | -28.3923
18.5386 | -35.2736 | -28.3997
18.5447 | -35.2741 | -28.4067
18.5508 | -35.2747 | -28.4137
18.557 | -35.2751 | -28.421
18.5631 | -35.2756 | -28.428
18.5692 | -35.2761 | -28.435
18.5753 | -35.2767 | -28.4419
18.5814 | -35.2772 | -28.4491
18.5875 | -35.2777 | -28.4565
18.5936 | -35.2781 | -28.4639
18.5997 | -35.2787 | -28.471
18.6058 | -35.2791 | -28.478
18.6119 | -35.2796 | -28.4852
18.618 | -35.2802 | -28.4924
18.6241 | -35.2807 | -28.4997
18.6302 | -35.2812 | -28.507
18.6363 | -35.2817 | -28.5139
18.6424 | -35.2823 | -28.5208
18.6485 | -35.2827 | -28.5279
18.6546 | -35.2833 | -28.5349
18.6607 | -35.2837 | -28.5419
18.6668 | -35.2843 | -28.5489
18.6729 | -35.2848 | -28.5561
18.679 | -35.2853 | -28.5628
18.6851 | -35.2858 | -28.5698
18.6912 | -35.2863 | -28.5768
18.6974 | -35.2868 | -28.584
18.7035 | -35.2873 | -28.591
18.7096 | -35.2878 | -28.5982
18.7157 | -35.2884 | -28.6051
18.7218 | -35.2889 | -28.6122
18.7279 | -35.2894 | -28.6191
18.734 | -35.2899 | -28.6262
18.7401 | -35.2904 | -28.633
18.7462 | -35.291 | -28.6395
18.7523 | -35.2915 | -28.6465
18.7584 | -35.292 | -28.6537
18.7645 | -35.2926 | -28.6607
18.7706 | -35.2931 | -28.6679
18.7767 | -35.2936 | -28.6748
18.7828 | -35.2941 | -28.6815
18.7889 | -35.2946 | -28.6884
18.795 | -35.2951 | -28.6954
18.8011 | -35.2956 | -28.7024
18.8072 | -35.2961 | -28.7095
18.8133 | -35.2966 | -28.7167
18.8194 | -35.2971 | -28.7234
18.8255 | -35.2976 | -28.7304
18.8316 | -35.2981 | -28.7375
18.8377 | -35.2986 | -28.7443
18.8439 | -35.2991 | -28.7512
18.85 | -35.2996 | -28.7582
18.8561 | -35.3001 | -28.7652
18.8622 | -35.3006 | -28.7724
18.8683 | -35.3011 | -28.7792
18.8744 | -35.3017 | -28.786
18.8805 | -35.3022 | -28.7931
18.8866 | -35.3028 | -28.8001
18.8927 | -35.3033 | -28.8069
18.8988 | -35.3038 | -28.8139
18.9049 | -35.3043 | -28.8209
18.911 | -35.3049 | -28.8276
18.9171 | -35.3054 | -28.8344
18.9232 | -35.3059 | -28.8414
18.9293 | -35.3064 | -28.8484
18.9354 | -35.3069 | -28.855
18.9415 | -35.3074 | -28.8619
18.9476 | -35.3079 | -28.869
18.9537 | -35.3084 | -28.8761
18.9598 | -35.309 | -28.8826
18.9659 | -35.3095 | -28.8896
18.972 | -35.31 | -28.8966
18.9781 | -35.3105 | -28.9035
18.9843 | -35.311 | -28.9105
18.9904 | -35.3115 | -28.9175
18.9965 | -35.312 | -28.9243
19.0026 | -35.3126 | -28.9314
19.0087 | -35.3131 | -28.9384
19.0148 | -35.3136 | -28.9452
19.0209 | -35.3141 | -28.9525
19.027 | -35.3147 | -28.9596
19.0331 | -35.3152 | -28.9665
19.0392 | -35.3157 | -28.9736
19.0453 | -35.3162 | -28.9805
19.0514 | -35.3167 | -28.9875
19.0575 | -35.3172 | -28.9945
19.0636 | -35.3177 | -29.0009
19.0697 | -35.3183 | -29.0082
19.0758 | -35.3188 | -29.0151
19.0819 | -35.3193 | -29.0223
19.088 | -35.3198 | -29.029
19.0941 | -35.3204 | -29.0357
19.1002 | -35.3209 | -29.0424
19.1063 | -35.3214 | -29.0494
19.1124 | -35.322 | -29.0563
19.1185 | -35.3225 | -29.0634
19.1246 | -35.323 | -29.0703
19.1308 | -35.3235 | -29.077
19.1369 | -35.324 | -29.0838
19.143 | -35.3245 | -29.0907
19.1491 | -35.3251 | -29.0976
19.1552 | -35.3256 | -29.1048
19.1613 | -35.3261 | -29.1117
19.1674 | -35.3266 | -29.1187
19.1735 | -35.3271 | -29.1256
19.1796 | -35.3277 | -29.1325
19.1857 | -35.3281 | -29.1395
19.1918 | -35.3286 | -29.1466
19.1979 | -35.3292 | -29.1536
19.204 | -35.3297 | -29.1604
19.2101 | -35.3302 | -29.1672
19.2162 | -35.3307 | -29.1744
19.2223 | -35.3312 | -29.181
19.2284 | -35.3317 | -29.1882
19.2345 | -35.3323 | -29.1953
19.2406 | -35.3328 | -29.2024
19.2467 | -35.3333 | -29.2094
19.2528 | -35.3338 | -29.2164
19.2589 | -35.3343 | -29.2234
19.265 | -35.3348 | -29.2303
19.2712 | -35.3353 | -29.2373
19.2773 | -35.3358 | -29.2442
19.2834 | -35.3363 | -29.251
19.2895 | -35.3369 | -29.2581
19.2956 | -35.3374 | -29.2653
19.3017 | -35.3379 | -29.2721
19.3078 | -35.3384 | -29.2791
19.3139 | -35.3389 | -29.2859
19.32 | -35.3395 | -29.2926
19.3261 | -35.34 | -29.2995
19.3322 | -35.3405 | -29.3064
19.3383 | -35.3411 | -29.3131
19.3444 | -35.3416 | -29.3198
19.3505 | -35.3421 | -29.3267
19.3566 | -35.3426 | -29.3335
19.3627 | -35.3431 | -29.3405
19.3688 | -35.3437 | -29.3475
19.3749 | -35.3442 | -29.3545
19.381 | -35.3447 | -29.361
19.3871 | -35.3452 | -29.3677
19.3932 | -35.3457 | -29.3747
19.3993 | -35.3462 | -29.3815
19.4054 | -35.3467 | -29.3886
19.4115 | -35.3473 | -29.3955
19.4177 | -35.3478 | -29.4022
19.4238 | -35.3483 | -29.409
19.4299 | -35.3488 | -29.4158
19.436 | -35.3494 | -29.4231
19.4421 | -35.3499 | -29.4296
19.4482 | -35.3504 | -29.4366
19.4543 | -35.351 | -29.4433
19.4604 | -35.3514 | -29.4501
19.4665 | -35.352 | -29.457
19.4726 | -35.3525 | -29.4636
19.4787 | -35.3531 | -29.4706
19.4848 | -35.3535 | -29.4774
19.4909 | -35.3541 | -29.4839
19.497 | -35.3546 | -29.4906
19.5031 | -35.3551 | -29.4974
19.5092 | -35.3556 | -29.5041
19.5153 | -35.3562 | -29.5111
19.5214 | -35.3567 | -29.5176
19.5275 | -35.3572 | -29.5243
19.5336 | -35.3577 | -29.5309
19.5397 | -35.3582 | -29.5377
19.5458 | -35.3588 | -29.5446
19.5519 | -35.3593 | -29.5513
19.5581 | -35.3598 | -29.558
19.5642 | -35.3604 | -29.5648
19.5703 | -35.3609 | -29.5714
19.5764 | -35.3614 | -29.5784
19.5825 | -35.3619 | -29.5852
19.5886 | -35.3623 | -29.592
19.5947 | -35.3629 | -29.5985
19.6008 | -35.3633 | -29.6052
19.6069 | -35.3639 | -29.6118
19.613 | -35.3644 | -29.6188
19.6191 | -35.3649 | -29.6256
19.6252 | -35.3654 | -29.6324
19.6313 | -35.3659 | -29.6389
19.6374 | -35.3664 | -29.6459
19.6435 | -35.3669 | -29.6523
19.6496 | -35.3675 | -29.6591
19.6557 | -35.368 | -29.6656
19.6618 | -35.3685 | -29.6721
19.6679 | -35.3691 | -29.679
19.674 | -35.3696 | -29.6857
19.6801 | -35.3701 | -29.6924
19.6862 | -35.3706 | -29.6989
19.6923 | -35.3712 | -29.7056
19.6984 | -35.3718 | -29.7122
19.7046 | -35.3723 | -29.719
19.7107 | -35.3728 | -29.7257
19.7168 | -35.3733 | -29.7328
19.7229 | -35.3738 | -29.7399
19.729 | -35.3743 | -29.7466
19.7351 | -35.3749 | -29.7531
19.7412 | -35.3755 | -29.7597
19.7473 | -35.376 | -29.7666
19.7534 | -35.3765 | -29.7736
19.7595 | -35.377 | -29.7805
19.7656 | -35.3775 | -29.7878
19.7717 | -35.378 | -29.7944
19.7778 | -35.3785 | -29.8014
19.7839 | -35.379 | -29.808
19.79 | -35.3795 | -29.8146
19.7961 | -35.38 | -29.8216
19.8022 | -35.3806 | -29.8285
19.8083 | -35.3811 | -29.8353
19.8144 | -35.3816 | -29.8419
19.8205 | -35.3821 | -29.8488
19.8266 | -35.3826 | -29.8558
19.8327 | -35.3831 | -29.8626
19.8388 | -35.3837 | -29.8695
19.845 | -35.3842 | -29.8763
19.8511 | -35.3847 | -29.8833
19.8572 | -35.3852 | -29.8901
19.8633 | -35.3857 | -29.8966
19.8694 | -35.3863 | -29.9034
19.8755 | -35.3868 | -29.9102
19.8816 | -35.3873 | -29.9172
19.8877 | -35.3879 | -29.9239
19.8938 | -35.3883 | -29.9304
19.8999 | -35.3889 | -29.9371
19.906 | -35.3894 | -29.9439
19.9121 | -35.3899 | -29.9506
19.9182 | -35.3904 | -29.9571
19.9243 | -35.3909 | -29.9642
19.9304 | -35.3914 | -29.9713
19.9365 | -35.3919 | -29.9781
19.9426 | -35.3925 | -29.9846
19.9487 | -35.393 | -29.9913
19.9548 | -35.3935 | -29.9984
19.9609 | -35.394 | -30.0049
19.967 | -35.3945 | -30.0116
19.9731 | -35.395 | -30.0184
19.9792 | -35.3955 | -30.0246
19.9853 | -35.396 | -30.0313
19.9915 | -35.3966 | -30.0382
19.9976 | -35.3971 | -30.0448
20.0037 | -35.3976 | -30.0515
20.0098 | -35.3981 | -30.0581
20.0159 | -35.3986 | -30.0649
20.022 | -35.3991 | -30.0718
20.0281 | -35.3997 | -30.0786
20.0342 | -35.4002 | -30.0852
20.0403 | -35.4007 | -30.0916
20.0464 | -35.4012 | -30.0987
20.0525 | -35.4018 | -30.1054
20.0586 | -35.4023 | -30.1119
20.0647 | -35.4028 | -30.1188
20.0708 | -35.4034 | -30.126
20.0769 | -35.4039 | -30.1326
20.083 | -35.4044 | -30.1391
20.0891 | -35.4049 | -30.1457
20.0952 | -35.4055 | -30.152
20.1013 | -35.4059 | -30.1588
20.1074 | -35.4065 | -30.1656
20.1135 | -35.407 | -30.1722
20.1196 | -35.4075 | -30.1789
20.1257 | -35.408 | -30.1856
20.1319 | -35.4085 | -30.1924
20.138 | -35.409 | -30.1991
20.1441 | -35.4095 | -30.2057
20.1502 | -35.4101 | -30.2125
20.1563 | -35.4106 | -30.2189
20.1624 | -35.4111 | -30.2257
20.1685 | -35.4116 | -30.2326
20.1746 | -35.4121 | -30.2393
20.1807 | -35.4126 | -30.2459
20.1868 | -35.4131 | -30.2526
20.1929 | -35.4137 | -30.259
20.199 | -35.4142 | -30.2659
20.2051 | -35.4147 | -30.2725
20.2112 | -35.4152 | -30.279
20.2173 | -35.4157 | -30.2858
20.2234 | -35.4162 | -30.2926
20.2295 | -35.4167 | -30.299
20.2356 | -35.4172 | -30.3059
20.2417 | -35.4177 | -30.3124
20.2478 | -35.4182 | -30.3193
20.2539 | -35.4187 | -30.3262
20.26 | -35.4193 | -30.333
20.2661 | -35.4198 | -30.3397
20.2723 | -35.4203 | -30.3463
20.2784 | -35.4208 | -30.3531
20.2845 | -35.4213 | -30.3596
20.2906 | -35.4219 | -30.3663
20.2967 | -35.4224 | -30.373
20.3028 | -35.4229 | -30.3796
20.3089 | -35.4234 | -30.3864
20.315 | -35.4239 | -30.393
20.3211 | -35.4244 | -30.3997
20.3272 | -35.425 | -30.4063
20.3333 | -35.4255 | -30.4127
20.3394 | -35.426 | -30.4191
20.3455 | -35.4265 | -30.4257
20.3516 | -35.4271 | -30.432
20.3577 | -35.4275 | -30.4384
20.3638 | -35.4281 | -30.4451
20.3699 | -35.4286 | -30.4517
20.376 | -35.4291 | -30.4584
20.3821 | -35.4296 | -30.465
20.3882 | -35.4301 | -30.4716
20.3943 | -35.4306 | -30.4785
20.4004 | -35.4312 | -30.4849
20.4065 | -35.4317 | -30.4913
20.4126 | -35.4322 | -30.4978
20.4188 | -35.4327 | -30.5047
20.4249 | -35.4333 | -30.5112
20.431 | -35.4337 | -30.518
20.4371 | -35.4344 | -30.5242
20.4432 | -35.4348 | -30.5309
20.4493 | -35.4353 | -30.5374
20.4554 | -35.4358 | -30.5439
20.4615 | -35.4364 | -30.5507
20.4676 | -35.4368 | -30.5573
20.4737 | -35.4374 | -30.5636
20.4798 | -35.4379 | -30.5703
20.4859 | -35.4384 | -30.5768
20.492 | -35.439 | -30.5836
20.4981 | -35.4395 | -30.5902
20.5042 | -35.44 | -30.5967
20.5103 | -35.4405 | -30.6032
20.5164 | -35.441 | -30.6099
20.5225 | -35.4415 | -30.6166
20.5286 | -35.442 | -30.6236
20.5347 | -35.4425 | -30.6301
20.5408 | -35.443 | -30.6366
20.5469 | -35.4436 | -30.643
20.553 | -35.4441 | -30.6494
20.5592 | -35.4446 | -30.6559
20.5653 | -35.445 | -30.6626
20.5714 | -35.4455 | -30.6686
20.5775 | -35.4461 | -30.6752
20.5836 | -35.4466 | -30.682
20.5897 | -35.4471 | -30.6884
20.5958 | -35.4476 | -30.6949
20.6019 | -35.4481 | -30.7016
20.608 | -35.4486 | -30.7083
20.6141 | -35.4491 | -30.7144
20.6202 | -35.4496 | -30.721
20.6263 | -35.4502 | -30.7273
20.6324 | -35.4506 | -30.7337
20.6385 | -35.4511 | -30.7401
20.6446 | -35.4517 | -30.7469
20.6507 | -35.4522 | -30.7533
20.6568 | -35.4527 | -30.7599
20.6629 | -35.4532 | -30.7663
20.669 | -35.4537 | -30.7729
20.6751 | -35.4542 | -30.7793
20.6812 | -35.4547 | -30.7854
20.6873 | -35.4552 | -30.7919
20.6934 | -35.4557 | -30.7985
20.6995 | -35.4562 | -30.8049
20.7057 | -35.4568 | -30.8114
20.7118 | -35.4573 | -30.8179
20.7179 | -35.4578 | -30.8239
20.724 | -35.4583 | -30.8304
20.7301 | -35.4588 | -30.837
20.7362 | -35.4594 | -30.8438
20.7423 | -35.4599 | -30.8501
20.7484 | -35.4604 | -30.8565
20.7545 | -35.4609 | -30.8628
20.7606 | -35.4615 | -30.8696
20.7667 | -35.462 | -30.8761
20.7728 | -35.4626 | -30.8827
20.7789 | -35.4631 | -30.8892
20.785 | -35.4636 | -30.8959
20.7911 | -35.4641 | -30.9025
20.7972 | -35.4646 | -30.909
20.8033 | -35.4651 | -30.9154
20.8094 | -35.4656 | -30.9222
20.8155 | -35.4661 | -30.9287
20.8216 | -35.4666 | -30.9354
20.8277 | -35.4671 | -30.9417
20.8338 | -35.4676 | -30.9482
20.8399 | -35.4681 | -30.9547
20.8461 | -35.4686 | -30.9611
20.8522 | -35.4691 | -30.9676
20.8583 | -35.4697 | -30.9741
20.8644 | -35.4702 | -30.9805
20.8705 | -35.4707 | -30.9871
20.8766 | -35.4712 | -30.9939
20.8827 | -35.4718 | -31.0003
20.8888 | -35.4723 | -31.0068
20.8949 | -35.4728 | -31.0131
20.901 | -35.4733 | -31.0196
20.9071 | -35.4738 | -31.0261
20.9132 | -35.4743 | -31.0326
20.9193 | -35.4748 | -31.0392
20.9254 | -35.4753 | -31.0455
20.9315 | -35.4759 | -31.0514
20.9376 | -35.4764 | -31.0579
20.9437 | -35.4769 | -31.0645
20.9498 | -35.4774 | -31.0709
20.9559 | -35.4779 | -31.0771
20.962 | -35.4785 | -31.0835
20.9681 | -35.479 | -31.09
20.9742 | -35.4795 | -31.0963
20.9803 | -35.48 | -31.1027
20.9864 | -35.4805 | -31.1093
20.9926 | -35.4811 | -31.1159
20.9987 | -35.4816 | -31.1222
21.0048 | -35.4821 | -31.1286
21.0109 | -35.4826 | -31.1349
21.017 | -35.4831 | -31.1412
21.0231 | -35.4836 | -31.1476
21.0292 | -35.4841 | -31.1542
21.0353 | -35.4846 | -31.1606
21.0414 | -35.4851 | -31.167
21.0475 | -35.4856 | -31.1733
21.0536 | -35.4861 | -31.1799
21.0597 | -35.4867 | -31.1861
21.0658 | -35.4871 | -31.1923
21.0719 | -35.4877 | -31.1989
21.078 | -35.4881 | -31.2055
21.0841 | -35.4887 | -31.2118
21.0902 | -35.4892 | -31.2183
21.0963 | -35.4896 | -31.225
21.1024 | -35.4901 | -31.2316
21.1085 | -35.4906 | -31.2378
21.1146 | -35.4912 | -31.2444
21.1207 | -35.4917 | -31.2506
21.1268 | -35.4922 | -31.2572
21.133 | -35.4927 | -31.2634
21.1391 | -35.4932 | -31.2697
21.1452 | -35.4937 | -31.2759
21.1513 | -35.4943 | -31.2822
21.1574 | -35.4948 | -31.2889
21.1635 | -35.4953 | -31.2951
21.1696 | -35.4958 | -31.3013
21.1757 | -35.4963 | -31.3073
21.1818 | -35.4968 | -31.3138
21.1879 | -35.4973 | -31.32
21.194 | -35.4978 | -31.3263
21.2001 | -35.4983 | -31.3323
21.2062 | -35.4988 | -31.3387
21.2123 | -35.4993 | -31.3448
21.2184 | -35.4998 | -31.3512
21.2245 | -35.5003 | -31.3577
21.2306 | -35.5008 | -31.3639
21.2367 | -35.5013 | -31.3704
21.2428 | -35.5019 | -31.3766
21.2489 | -35.5024 | -31.3826
21.255 | -35.5029 | -31.389
21.2611 | -35.5034 | -31.3953
21.2672 | -35.5039 | -31.4018
21.2733 | -35.5044 | -31.4082
21.2795 | -35.5049 | -31.4142
21.2856 | -35.5055 | -31.4206
21.2917 | -35.506 | -31.427
21.2978 | -35.5065 | -31.4333
21.3039 | -35.507 | -31.4395
21.31 | -35.5075 | -31.4458
21.3161 | -35.508 | -31.452
21.3222 | -35.5086 | -31.4581
21.3283 | -35.509 | -31.4645
21.3344 | -35.5095 | -31.4709
21.3405 | -35.51 | -31.4773
21.3466 | -35.5106 | -31.4834
21.3527 | -35.5111 | -31.4894
21.3588 | -35.5117 | -31.4956
21.3649 | -35.5122 | -31.5018
21.371 | -35.5127 | -31.508
21.3771 | -35.5132 | -31.5141
21.3832 | -35.5137 | -31.5204
21.3893 | -35.5142 | -31.5263
21.3954 | -35.5147 | -31.5325
21.4015 | -35.5152 | -31.5387
21.4076 | -35.5157 | -31.545
21.4137 | -35.5163 | -31.5514
21.4199 | -35.5167 | -31.5577
21.426 | -35.5172 | -31.564
21.4321 | -35.5177 | -31.5702
21.4382 | -35.5182 | -31.5763
21.4443 | -35.5187 | -31.5824
21.4504 | -35.5192 | -31.5886
21.4565 | -35.5197 | -31.5947
21.4626 | -35.5203 | -31.601
21.4687 | -35.5208 | -31.6071
21.4748 | -35.5213 | -31.6133
21.4809 | -35.5218 | -31.6197
21.487 | -35.5223 | -31.6258
21.4931 | -35.5228 | -31.6316
21.4992 | -35.5233 | -31.6375
21.5053 | -35.5237 | -31.6438
21.5114 | -35.5243 | -31.6499
21.5175 | -35.5247 | -31.6561
21.5236 | -35.5252 | -31.6624
21.5297 | -35.5257 | -31.6682
21.5358 | -35.5262 | -31.6741
21.5419 | -35.5267 | -31.6807
21.548 | -35.5272 | -31.6869
21.5541 | -35.5277 | -31.6932
21.5602 | -35.5282 | -31.6994
21.5664 | -35.5287 | -31.7055
21.5725 | -35.5292 | -31.7118
21.5786 | -35.5296 | -31.7177
21.5847 | -35.5301 | -31.7241
21.5908 | -35.5306 | -31.7303
21.5969 | -35.5311 | -31.7363
21.603 | -35.5316 | -31.7425
21.6091 | -35.5321 | -31.749
21.6152 | -35.5327 | -31.7553
21.6213 | -35.5332 | -31.7617
21.6274 | -35.5337 | -31.7678
21.6335 | -35.5342 | -31.774
21.6396 | -35.5347 | -31.7802
21.6457 | -35.5352 | -31.7868
21.6518 | -35.5357 | -31.793
21.6579 | -35.5362 | -31.7988
21.664 | -35.5367 | -31.8053
21.6701 | -35.5372 | -31.811
21.6762 | -35.5377 | -31.8172
21.6823 | -35.5382 | -31.8232
21.6884 | -35.5387 | -31.8295
21.6945 | -35.5392 | -31.8359
21.7006 | -35.5396 | -31.8422
21.7068 | -35.5402 | -31.8478
21.7129 | -35.5407 | -31.854
21.719 | -35.5412 | -31.8602
21.7251 | -35.5417 | -31.8664
21.7312 | -35.5422 | -31.8726
21.7373 | -35.5427 | -31.8785
21.7434 | -35.5432 | -31.8851
21.7495 | -35.5437 | -31.8911
21.7556 | -35.5442 | -31.8976
21.7617 | -35.5447 | -31.9038
21.7678 | -35.5452 | -31.91
21.7739 | -35.5457 | -31.916
21.78 | -35.5462 | -31.9221
21.7861 | -35.5468 | -31.9282
21.7922 | -35.5472 | -31.9344
21.7983 | -35.5478 | -31.9405
21.8044 | -35.5483 | -31.9467
21.8105 | -35.5488 | -31.9527
21.8166 | -35.5493 | -31.959
21.8227 | -35.5498 | -31.9651
21.8288 | -35.5503 | -31.9711
21.8349 | -35.5508 | -31.9774
21.841 | -35.5513 | -31.9838
21.8471 | -35.5518 | -31.9898
21.8533 | -35.5523 | -31.9956
21.8594 | -35.5528 | -32.0021
21.8655 | -35.5533 | -32.0084
21.8716 | -35.5538 | -32.0146
21.8777 | -35.5543 | -32.0208
21.8838 | -35.5549 | -32.0268
21.8899 | -35.5554 | -32.0328
21.896 | -35.5559 | -32.0392
21.9021 | -35.5564 | -32.0454
21.9082 | -35.5569 | -32.0516
21.9143 | -35.5574 | -32.0574
21.9204 | -35.5579 | -32.0635
21.9265 | -35.5583 | -32.0695
21.9326 | -35.5589 | -32.076
21.9387 | -35.5594 | -32.0819
21.9448 | -35.5599 | -32.0881
21.9509 | -35.5604 | -32.0941
21.957 | -35.5609 | -32.1001
21.9631 | -35.5614 | -32.1064
21.9692 | -35.5619 | -32.1127
21.9753 | -35.5624 | -32.1188
21.9814 | -35.5629 | -32.1247
21.9875 | -35.5634 | -32.1307
21.9937 | -35.5639 | -32.137
21.9998 | -35.5644 | -32.1431
22.0059 | -35.5649 | -32.1495
22.012 | -35.5654 | -32.1555
22.0181 | -35.5659 | -32.1614
22.0242 | -35.5664 | -32.1674
22.0303 | -35.5669 | -32.1737
22.0364 | -35.5674 | -32.1796
22.0425 | -35.5679 | -32.1857
22.0486 | -35.5683 | -32.1918
22.0547 | -35.5688 | -32.1977
22.0608 | -35.5693 | -32.2037
22.0669 | -35.5698 | -32.2097
22.073 | -35.5703 | -32.2156
22.0791 | -35.5708 | -32.2217
22.0852 | -35.5713 | -32.2278
22.0913 | -35.5718 | -32.2336
22.0974 | -35.5722 | -32.2395
22.1035 | -35.5727 | -32.2455
22.1096 | -35.5732 | -32.2517
22.1157 | -35.5737 | -32.2577
22.1218 | -35.5742 | -32.2641
22.1279 | -35.5747 | -32.27
22.134 | -35.5752 | -32.276
22.1402 | -35.5757 | -32.2819
22.1463 | -35.5761 | -32.288
22.1524 | -35.5766 | -32.2942
22.1585 | -35.5771 | -32.3004
22.1646 | -35.5776 | -32.3058
22.1707 | -35.5781 | -32.3121
22.1768 | -35.5786 | -32.3183
22.1829 | -35.5791 | -32.3244
22.189 | -35.5796 | -32.3305
22.1951 | -35.5801 | -32.3361
22.2012 | -35.5806 | -32.3422
22.2073 | -35.5811 | -32.3483
22.2134 | -35.5816 | -32.3543
22.2195 | -35.5821 | -32.3604
22.2256 | -35.5826 | -32.3664
22.2317 | -35.5831 | -32.3724
22.2378 | -35.5836 | -32.3782
22.2439 | -35.5841 | -32.3844
22.25 | -35.5846 | -32.3905
22.2561 | -35.5852 | -32.3966
22.2622 | -35.5856 | -32.4028
22.2683 | -35.5861 | -32.4085
22.2744 | -35.5866 | -32.4144
22.2806 | -35.5871 | -32.4204
22.2867 | -35.5875 | -32.4263
22.2928 | -35.588 | -32.4322
22.2989 | -35.5885 | -32.4381
22.305 | -35.589 | -32.444
22.3111 | -35.5895 | -32.4502
22.3172 | -35.59 | -32.4563
22.3233 | -35.5905 | -32.4624
22.3294 | -35.591 | -32.4684
22.3355 | -35.5916 | -32.4742
22.3416 | -35.5921 | -32.4802
22.3477 | -35.5926 | -32.4864
22.3538 | -35.5931 | -32.4922
22.3599 | -35.5935 | -32.4982
22.366 | -35.594 | -32.5042
22.3721 | -35.5945 | -32.5099
22.3782 | -35.595 | -32.5161
22.3843 | -35.5955 | -32.5222
22.3904 | -35.596 | -32.5285
22.3965 | -35.5965 | -32.5345
22.4026 | -35.597 | -32.5404
22.4087 | -35.5975 | -32.5463
22.4148 | -35.598 | -32.5523
22.4209 | -35.5984 | -32.5584
22.4271 | -35.5989 | -32.5643
22.4332 | -35.5994 | -32.57
22.4393 | -35.5999 | -32.5761
22.4454 | -35.6004 | -32.5816
22.4515 | -35.6008 | -32.5877
22.4576 | -35.6014 | -32.5937
22.4637 | -35.6018 | -32.5999
22.4698 | -35.6023 | -32.6059
22.4759 | -35.6028 | -32.6116
22.482 | -35.6033 | -32.6176
22.4881 | -35.6038 | -32.6237
22.4942 | -35.6043 | -32.6297
22.5003 | -35.6048 | -32.6357
22.5064 | -35.6053 | -32.6416
22.5125 | -35.6057 | -32.6474
22.5186 | -35.6062 | -32.6532
22.5247 | -35.6067 | -32.6593
22.5308 | -35.6072 | -32.6655
22.5369 | -35.6077 | -32.6718
22.543 | -35.6082 | -32.6777
22.5491 | -35.6087 | -32.6832
22.5552 | -35.6091 | -32.6893
22.5613 | -35.6096 | -32.6953
22.5675 | -35.6101 | -32.7014
22.5736 | -35.6105 | -32.7075
22.5797 | -35.611 | -32.7135
22.5858 | -35.6115 | -32.7193
22.5919 | -35.612 | -32.7251
22.598 | -35.6125 | -32.7313
22.6041 | -35.6129 | -32.7372
22.6102 | -35.6134 | -32.7432
22.6163 | -35.6139 | -32.7489
22.6224 | -35.6144 | -32.7549
22.6285 | -35.6149 | -32.7608
22.6346 | -35.6153 | -32.7666
22.6407 | -35.6158 | -32.7728
22.6468 | -35.6163 | -32.7788
22.6529 | -35.6168 | -32.7845
22.659 | -35.6173 | -32.7904
22.6651 | -35.6177 | -32.7964
22.6712 | -35.6182 | -32.8022
22.6773 | -35.6187 | -32.808
22.6834 | -35.6192 | -32.8141
22.6895 | -35.6197 | -32.8198
22.6956 | -35.6202 | -32.8257
22.7017 | -35.6207 | -32.8317
22.7079 | -35.6211 | -32.8379
22.714 | -35.6216 | -32.8438
22.7201 | -35.6221 | -32.8497
22.7262 | -35.6226 | -32.8556
22.7323 | -35.6231 | -32.8614
22.7384 | -35.6236 | -32.8674
22.7445 | -35.624 | -32.8735
22.7506 | -35.6245 | -32.8793
22.7567 | -35.625 | -32.8853
22.7628 | -35.6255 | -32.891
22.7689 | -35.626 | -32.8967
22.775 | -35.6265 | -32.9029
22.7811 | -35.627 | -32.9089
22.7872 | -35.6274 | -32.9148
22.7933 | -35.6279 | -32.9205
22.7994 | -35.6284 | -32.9264
22.8055 | -35.6289 | -32.9323
22.8116 | -35.6294 | -32.9381
22.8177 | -35.6298 | -32.9441
22.8238 | -35.6303 | -32.9502
22.8299 | -35.6308 | -32.9559
22.836 | -35.6313 | -32.962
22.8421 | -35.6318 | -32.9679
22.8482 | -35.6322 | -32.9738
22.8544 | -35.6327 | -32.9799
22.8605 | -35.6332 | -32.9858
22.8666 | -35.6337 | -32.9916
22.8727 | -35.6342 | -32.9976
22.8788 | -35.6346 | -33.0033
22.8849 | -35.6351 | -33.0092
22.891 | -35.6356 | -33.0152
22.8971 | -35.6361 | -33.0211
22.9032 | -35.6366 | -33.0266
22.9093 | -35.637 | -33.0325
22.9154 | -35.6375 | -33.0384
22.9215 | -35.638 | -33.0443
22.9276 | -35.6384 | -33.0502
22.9337 | -35.6389 | -33.0557
22.9398 | -35.6394 | -33.0617
22.9459 | -35.6399 | -33.0676
22.952 | -35.6403 | -33.0736
22.9581 | -35.6408 | -33.0795
22.9642 | -35.6413 | -33.0855
22.9703 | -35.6417 | -33.0911
22.9764 | -35.6422 | -33.0969
22.9825 | -35.6427 | -33.1028
22.9886 | -35.6432 | -33.1087
22.9948 | -35.6436 | -33.1146
23.0009 | -35.6441 | -33.1205
23.007 | -35.6446 | -33.1264
23.0131 | -35.6451 | -33.1322
23.0192 | -35.6455 | -33.138
23.0253 | -35.646 | -33.144
23.0314 | -35.6465 | -33.1498
23.0375 | -35.647 | -33.1557
23.0436 | -35.6474 | -33.1614
23.0497 | -35.6479 | -33.1673
23.0558 | -35.6484 | -33.173
23.0619 | -35.6489 | -33.1789
23.068 | -35.6494 | -33.1849
23.0741 | -35.6499 | -33.1903
23.0802 | -35.6503 | -33.196
23.0863 | -35.6508 | -33.2019
23.0924 | -35.6512 | -33.2078
23.0985 | -35.6517 | -33.2136
23.1046 | -35.6522 | -33.2195
23.1107 | -35.6526 | -33.2253
23.1168 | -35.6531 | -33.2313
23.1229 | -35.6535 | -33.2371
23.129 | -35.654 | -33.2432
23.1351 | -35.6545 | -33.2491
23.1413 | -35.655 | -33.255
23.1474 | -35.6554 | -33.2607
23.1535 | -35.6559 | -33.2665
23.1596 | -35.6564 | -33.2721
23.1657 | -35.6568 | -33.2778
23.1718 | -35.6573 | -33.2837
23.1779 | -35.6577 | -33.2898
23.184 | -35.6582 | -33.2953
23.1901 | -35.6587 | -33.3012
23.1962 | -35.6592 | -33.3069
23.2023 | -35.6596 | -33.3126
23.2084 | -35.6601 | -33.3184
23.2145 | -35.6606 | -33.3241
23.2206 | -35.6611 | -33.3299
23.2267 | -35.6615 | -33.3358
23.2328 | -35.662 | -33.3415
23.2389 | -35.6625 | -33.3475
23.245 | -35.663 | -33.3532
23.2511 | -35.6634 | -33.3587
23.2572 | -35.6639 | -33.3644
23.2633 | -35.6644 | -33.3701
23.2694 | -35.6649 | -33.3762
23.2755 | -35.6653 | -33.3822
23.2817 | -35.6658 | -33.3879
23.2878 | -35.6662 | -33.3935
23.2939 | -35.6668 | -33.3994
23.3 | -35.6672 | -33.4054
23.3061 | -35.6677 | -33.4112
23.3122 | -35.6682 | -33.4169
23.3183 | -35.6686 | -33.4227
23.3244 | -35.6691 | -33.4284
23.3305 | -35.6695 | -33.4343
23.3366 | -35.6699 | -33.4404
23.3427 | -35.6704 | -33.4459
23.3488 | -35.6709 | -33.4518
23.3549 | -35.6713 | -33.4577
23.361 | -35.6718 | -33.4635
23.3671 | -35.6723 | -33.4694
23.3732 | -35.6728 | -33.475
23.3793 | -35.6733 | -33.4807
23.3854 | -35.6737 | -33.4866
23.3915 | -35.6742 | -33.4924
23.3976 | -35.6747 | -33.4983
23.4037 | -35.6751 | -33.5038
23.4098 | -35.6756 | -33.5094
23.4159 | -35.6761 | -33.5152
23.422 | -35.6766 | -33.521
23.4282 | -35.677 | -33.5264
23.4343 | -35.6775 | -33.5321
23.4404 | -35.6779 | -33.5381
23.4465 | -35.6784 | -33.5438
23.4526 | -35.6789 | -33.5498
23.4587 | -35.6793 | -33.5556
23.4648 | -35.6797 | -33.5613
23.4709 | -35.6802 | -33.567
23.477 | -35.6806 | -33.573
23.4831 | -35.6811 | -33.5787
23.4892 | -35.6815 | -33.5848
23.4953 | -35.682 | -33.5905
23.5014 | -35.6824 | -33.5962
23.5075 | -35.6829 | -33.602
23.5136 | -35.6833 | -33.6077
23.5197 | -35.6838 | -33.6135
23.5258 | -35.6843 | -33.6193
23.5319 | -35.6847 | -33.6248
23.538 | -35.6852 | -33.6306
23.5441 | -35.6856 | -33.6364
23.5502 | -35.686 | -33.6423
23.5563 | -35.6865 | -33.6479
23.5624 | -35.687 | -33.6538
23.5686 | -35.6874 | -33.6595
23.5747 | -35.6879 | -33.665
23.5808 | -35.6883 | -33.6707
23.5869 | -35.6888 | -33.6764
23.593 | -35.6893 | -33.6822
23.5991 | -35.6897 | -33.6879
23.6052 | -35.6902 | -33.6934
23.6113 | -35.6907 | -33.6992
23.6174 | -35.6911 | -33.7049
23.6235 | -35.6916 | -33.7107
23.6296 | -35.692 | -33.7167
23.6357 | -35.6924 | -33.7227
23.6418 | -35.6929 | -33.7281
23.6479 | -35.6934 | -33.7341
23.654 | -35.6938 | -33.7398
23.6601 | -35.6943 | -33.7455
23.6662 | -35.6947 | -33.7513
23.6723 | -35.6952 | -33.7569
23.6784 | -35.6957 | -33.7626
23.6845 | -35.6961 | -33.7685
23.6906 | -35.6966 | -33.7744
23.6967 | -35.697 | -33.78
23.7028 | -35.6975 | -33.7858
23.7089 | -35.6979 | -33.7918
23.7151 | -35.6984 | -33.7975
23.7212 | -35.6988 | -33.8033
23.7273 | -35.6992 | -33.8092
23.7334 | -35.6997 | -33.8149
23.7395 | -35.7002 | -33.8205
23.7456 | -35.7006 | -33.8262
23.7517 | -35.701 | -33.8319
23.7578 | -35.7015 | -33.8375
23.7639 | -35.702 | -33.8433
23.77 | -35.7024 | -33.849
23.7761 | -35.7029 | -33.8545
23.7822 | -35.7033 | -33.8604
23.7883 | -35.7038 | -33.8661
23.7944 | -35.7042 | -33.872
23.8005 | -35.7047 | -33.8779
23.8066 | -35.7052 | -33.8836
23.8127 | -35.7056 | -33.8895
23.8188 | -35.7061 | -33.8953
23.8249 | -35.7065 | -33.9011
23.831 | -35.707 | -33.907
23.8371 | -35.7074 | -33.9125
23.8432 | -35.7079 | -33.9181
23.8493 | -35.7083 | -33.9237
23.8555 | -35.7088 | -33.9295
23.8616 | -35.7092 | -33.9354
23.8677 | -35.7097 | -33.9408
23.8738 | -35.7101 | -33.9466
23.8799 | -35.7106 | -33.9523
23.886 | -35.711 | -33.9579
23.8921 | -35.7114 | -33.9637
23.8982 | -35.7119 | -33.9696
23.9043 | -35.7123 | -33.9754
23.9104 | -35.7128 | -33.9812
23.9165 | -35.7132 | -33.9868
23.9226 | -35.7136 | -33.9923
23.9287 | -35.7141 | -33.998
23.9348 | -35.7145 | -34.0035
23.9409 | -35.715 | -34.0094
23.947 | -35.7154 | -34.015
23.9531 | -35.7158 | -34.0205
23.9592 | -35.7163 | -34.0263
23.9653 | -35.7167 | -34.0321
23.9714 | -35.7172 | -34.0377
23.9775 | -35.7176 | -34.0434
23.9836 | -35.7181 | -34.0492
23.9897 | -35.7186 | -34.0549
23.9958 | -35.719 | -34.0607
24.002 | -35.7194 | -34.0663
24.0081 | -35.7199 | -34.0718
24.0142 | -35.7203 | -34.0776
24.0203 | -35.7208 | -34.0833
24.0264 | -35.7212 | -34.0891
24.0325 | -35.7216 | -34.0948
24.0386 | -35.722 | -34.1006
24.0447 | -35.7225 | -34.1061
24.0508 | -35.723 | -34.1119
24.0569 | -35.7234 | -34.1178
24.063 | -35.7239 | -34.1233
24.0691 | -35.7243 | -34.1288
24.0752 | -35.7248 | -34.1346
24.0813 | -35.7252 | -34.1402
24.0874 | -35.7257 | -34.1459
24.0935 | -35.7261 | -34.1516
24.0996 | -35.7265 | -34.1569
24.1057 | -35.7269 | -34.1625
24.1118 | -35.7273 | -34.1683
24.1179 | -35.7277 | -34.1742
24.124 | -35.7282 | -34.1797
24.1301 | -35.7286 | -34.1851
24.1362 | -35.7291 | -34.1909
24.1424 | -35.7295 | -34.1965
24.1485 | -35.7299 | -34.2019
24.1546 | -35.7304 | -34.2078
24.1607 | -35.7308 | -34.2134
24.1668 | -35.7312 | -34.2189
24.1729 | -35.7316 | -34.2246
24.179 | -35.732 | -34.2301
24.1851 | -35.7325 | -34.2359
24.1912 | -35.7329 | -34.2417
24.1973 | -35.7334 | -34.2475
24.2034 | -35.7338 | -34.2528
24.2095 | -35.7343 | -34.2587
24.2156 | -35.7347 | -34.2646
24.2217 | -35.7351 | -34.2702
24.2278 | -35.7356 | -34.2758
24.2339 | -35.736 | -34.2813
24.24 | -35.7364 | -34.2871
24.2461 | -35.7369 | -34.2927
24.2522 | -35.7373 | -34.2984
24.2583 | -35.7378 | -34.3042
24.2644 | -35.7382 | -34.31
24.2705 | -35.7386 | -34.3159
24.2766 | -35.7391 | -34.3215
24.2827 | -35.7395 | -34.327
24.2889 | -35.7399 | -34.3326
24.295 | -35.7404 | -34.3384
24.3011 | -35.7408 | -34.3441
24.3072 | -35.7412 | -34.3494
24.3133 | -35.7416 | -34.355
24.3194 | -35.742 | -34.3609
24.3255 | -35.7425 | -34.3667
24.3316 | -35.7429 | -34.3724
24.3377 | -35.7433 | -34.378
24.3438 | -35.7437 | -34.3837
24.3499 | -35.7441 | -34.3893
24.356 | -35.7445 | -34.395
24.3621 | -35.7449 | -34.4006
24.3682 | -35.7453 | -34.4061
24.3743 | -35.7457 | -34.4119
24.3804 | -35.7461 | -34.4171
24.3865 | -35.7465 | -34.423
24.3926 | -35.7469 | -34.4287
24.3987 | -35.7473 | -34.4346
24.4048 | -35.7477 | -34.4402
24.4109 | -35.7481 | -34.4457
24.417 | -35.7486 | -34.4513
24.4231 | -35.749 | -34.457
24.4293 | -35.7495 | -34.4626
24.4354 | -35.7499 | -34.4681
24.4415 | -35.7503 | -34.4736
24.4476 | -35.7507 | -34.479
24.4537 | -35.7512 | -34.4844
24.4598 | -35.7516 | -34.4903
24.4659 | -35.752 | -34.4961
24.472 | -35.7525 | -34.5017
24.4781 | -35.7529 | -34.5075
24.4842 | -35.7533 | -34.5129
24.4903 | -35.7537 | -34.5184
24.4964 | -35.7541 | -34.5239
24.5025 | -35.7545 | -34.5298
24.5086 | -35.7549 | -34.5354
24.5147 | -35.7554 | -34.5408
24.5208 | -35.7558 | -34.5462
24.5269 | -35.7562 | -34.5516
24.533 | -35.7567 | -34.557
24.5391 | -35.7571 | -34.5628
24.5452 | -35.7576 | -34.5684
24.5513 | -35.758 | -34.574
24.5574 | -35.7584 | -34.5797
24.5635 | -35.7588 | -34.5852
24.5696 | -35.7593 | -34.5908
24.5758 | -35.7597 | -34.5964
24.5819 | -35.7601 | -34.6021
24.588 | -35.7605 | -34.6076
24.5941 | -35.761 | -34.6131
24.6002 | -35.7614 | -34.6188
24.6063 | -35.7618 | -34.6245
24.6124 | -35.7622 | -34.6303
24.6185 | -35.7626 | -34.6361
24.6246 | -35.763 | -34.6415
24.6307 | -35.7634 | -34.6473
24.6368 | -35.7638 | -34.653
24.6429 | -35.7642 | -34.6584
24.649 | -35.7647 | -34.6639
24.6551 | -35.7651 | -34.6694
24.6612 | -35.7655 | -34.6748
24.6673 | -35.7659 | -34.6805
24.6734 | -35.7663 | -34.6861
24.6795 | -35.7668 | -34.6918
24.6856 | -35.7672 | -34.6975
24.6917 | -35.7676 | -34.703
24.6978 | -35.768 | -34.7085
24.7039 | -35.7684 | -34.714
24.71 | -35.7689 | -34.7197
24.7162 | -35.7693 | -34.7253
24.7223 | -35.7697 | -34.7308
24.7284 | -35.7701 | -34.7361
24.7345 | -35.7705 | -34.7417
24.7406 | -35.771 | -34.7474
24.7467 | -35.7714 | -34.7531
24.7528 | -35.7718 | -34.7586
24.7589 | -35.7722 | -34.7638
24.765 | -35.7726 | -34.7688
24.7711 | -35.773 | -34.7745
24.7772 | -35.7734 | -34.7802
24.7833 | -35.7738 | -34.7858
24.7894 | -35.7742 | -34.7914
24.7955 | -35.7746 | -34.7971
24.8016 | -35.775 | -34.8026
24.8077 | -35.7755 | -34.8083
24.8138 | -35.7759 | -34.814
24.8199 | -35.7763 | -34.8196
24.826 | -35.7767 | -34.825
24.8321 | -35.7772 | -34.8306
24.8382 | -35.7776 | -34.836
24.8443 | -35.778 | -34.8416
24.8504 | -35.7784 | -34.847
24.8565 | -35.7788 | -34.8526
24.8627 | -35.7792 | -34.8582
24.8688 | -35.7796 | -34.8637
24.8749 | -35.7801 | -34.8691
24.881 | -35.7805 | -34.8746
24.8871 | -35.7809 | -34.8802
24.8932 | -35.7813 | -34.8857
24.8993 | -35.7818 | -34.8911
24.9054 | -35.7822 | -34.8966
24.9115 | -35.7826 | -34.9022
24.9176 | -35.783 | -34.9078
24.9237 | -35.7834 | -34.9134
24.9298 | -35.7838 | -34.9189
24.9359 | -35.7843 | -34.9245
24.942 | -35.7847 | -34.9299
24.9481 | -35.7851 | -34.9355
24.9542 | -35.7855 | -34.9412
24.9603 | -35.7859 | -34.9467
24.9664 | -35.7863 | -34.9522
24.9725 | -35.7867 | -34.9579
24.9786 | -35.787 | -34.9633
24.9847 | -35.7874 | -34.9687
24.9908 | -35.7878 | -34.9744
24.9969 | -35.7882 | -34.98
25.0031 | -35.7886 | -34.9854
25.0092 | -35.789 | -34.9909
25.0153 | -35.7894 | -34.9963
25.0214 | -35.7898 | -35.002
25.0275 | -35.7902 | -35.0075
25.0336 | -35.7907 | -35.0132
25.0397 | -35.7911 | -35.0189
25.0458 | -35.7915 | -35.0243
25.0519 | -35.7919 | -35.03
25.058 | -35.7923 | -35.0355
25.0641 | -35.7927 | -35.041
25.0702 | -35.7931 | -35.0464
25.0763 | -35.7935 | -35.0518
25.0824 | -35.7939 | -35.0574
25.0885 | -35.7942 | -35.0628
25.0946 | -35.7947 | -35.0684
25.1007 | -35.7951 | -35.074
25.1068 | -35.7955 | -35.0794
25.1129 | -35.7959 | -35.0849
25.119 | -35.7963 | -35.0902
25.1251 | -35.7967 | -35.0957
25.1312 | -35.7971 | -35.1013
25.1373 | -35.7975 | -35.1069
25.1435 | -35.7979 | -35.1126
25.1496 | -35.7983 | -35.118
25.1557 | -35.7987 | -35.1235
25.1618 | -35.7991 | -35.1292
25.1679 | -35.7995 | -35.1346
25.174 | -35.7999 | -35.1401
25.1801 | -35.8003 | -35.1456
25.1862 | -35.8007 | -35.151
25.1923 | -35.8011 | -35.1565
25.1984 | -35.8015 | -35.1621
25.2045 | -35.802 | -35.1676
25.2106 | -35.8023 | -35.1731
25.2167 | -35.8027 | -35.1787
25.2228 | -35.8031 | -35.1841
25.2289 | -35.8035 | -35.1895
25.235 | -35.8039 | -35.195
25.2411 | -35.8043 | -35.2005
25.2472 | -35.8047 | -35.2059
25.2533 | -35.8051 | -35.2115
25.2594 | -35.8054 | -35.2167
25.2655 | -35.8058 | -35.2222
25.2716 | -35.8062 | -35.2277
25.2777 | -35.8066 | -35.2334
25.2838 | -35.807 | -35.2391
25.29 | -35.8074 | -35.2445
25.2961 | -35.8078 | -35.25
25.3022 | -35.8082 | -35.2554
25.3083 | -35.8086 | -35.2609
25.3144 | -35.809 | -35.2662
25.3205 | -35.8093 | -35.2717
25.3266 | -35.8097 | -35.2771
25.3327 | -35.8101 | -35.2826
25.3388 | -35.8105 | -35.288
25.3449 | -35.8109 | -35.2935
25.351 | -35.8113 | -35.2989
25.3571 | -35.8116 | -35.3045
25.3632 | -35.812 | -35.3097
25.3693 | -35.8124 | -35.3152
25.3754 | -35.8128 | -35.3207
25.3815 | -35.8132 | -35.3261
25.3876 | -35.8136 | -35.3317
25.3937 | -35.814 | -35.3371
25.3998 | -35.8144 | -35.3424
25.4059 | -35.8148 | -35.3479
25.412 | -35.8152 | -35.3533
25.4181 | -35.8156 | -35.3589
25.4242 | -35.8159 | -35.3645
25.4304 | -35.8163 | -35.3699
25.4365 | -35.8167 | -35.3754
25.4426 | -35.8171 | -35.381
25.4487 | -35.8175 | -35.3866
25.4548 | -35.8179 | -35.3921
25.4609 | -35.8183 | -35.3976
25.467 | -35.8186 | -35.4029
25.4731 | -35.8191 | -35.4084
25.4792 | -35.8194 | -35.4138
25.4853 | -35.8199 | -35.4191
25.4914 | -35.8203 | -35.4246
25.4975 | -35.8207 | -35.4299
25.5036 | -35.821 | -35.4353
25.5097 | -35.8214 | -35.4406
25.5158 | -35.8218 | -35.4461
25.5219 | -35.8222 | -35.4516
25.528 | -35.8226 | -35.4571
25.5341 | -35.8231 | -35.4625
25.5402 | -35.8235 | -35.468
25.5463 | -35.8239 | -35.4738
25.5524 | -35.8243 | -35.4794
25.5585 | -35.8247 | -35.4848
25.5646 | -35.8251 | -35.4905
25.5707 | -35.8254 | -35.4959
25.5769 | -35.8258 | -35.5014
25.583 | -35.8262 | -35.5071
25.5891 | -35.8266 | -35.5126
25.5952 | -35.827 | -35.518
25.6013 | -35.8274 | -35.5234
25.6074 | -35.8278 | -35.5286
25.6135 | -35.8282 | -35.534
25.6196 | -35.8287 | -35.5397
25.6257 | -35.829 | -35.545
25.6318 | -35.8294 | -35.5503
25.6379 | -35.8298 | -35.5557
25.644 | -35.8302 | -35.5608
25.6501 | -35.8305 | -35.5659
25.6562 | -35.8309 | -35.5714
25.6623 | -35.8313 | -35.5768
25.6684 | -35.8317 | -35.5824
25.6745 | -35.8321 | -35.5879
25.6806 | -35.8325 | -35.5931
25.6867 | -35.8329 | -35.5986
25.6928 | -35.8333 | -35.604
25.6989 | -35.8336 | -35.6096
25.705 | -35.834 | -35.6152
25.7111 | -35.8344 | -35.6208
25.7173 | -35.8347 | -35.6261
25.7234 | -35.8351 | -35.6315
25.7295 | -35.8355 | -35.637
25.7356 | -35.8359 | -35.6427
25.7417 | -35.8362 | -35.648
25.7478 | -35.8366 | -35.6533
25.7539 | -35.837 | -35.6585
25.76 | -35.8374 | -35.664
25.7661 | -35.8378 | -35.6693
25.7722 | -35.8382 | -35.6747
25.7783 | -35.8385 | -35.6803
25.7844 | -35.8389 | -35.6854
25.7905 | -35.8393 | -35.691
25.7966 | -35.8396 | -35.6963
25.8027 | -35.84 | -35.7018
25.8088 | -35.8403 | -35.7074
25.8149 | -35.8407 | -35.7128
25.821 | -35.8411 | -35.7181
25.8271 | -35.8415 | -35.7236
25.8332 | -35.8419 | -35.7291
25.8393 | -35.8423 | -35.7346
25.8454 | -35.8426 | -35.7398
25.8515 | -35.8431 | -35.7453
25.8576 | -35.8434 | -35.7506
25.8638 | -35.8438 | -35.7562
25.8699 | -35.8442 | -35.7617
25.876 | -35.8446 | -35.7672
25.8821 | -35.845 | -35.7726
25.8882 | -35.8454 | -35.7777
25.8943 | -35.8458 | -35.7833
25.9004 | -35.8461 | -35.7889
25.9065 | -35.8466 | -35.7944
25.9126 | -35.847 | -35.7998
25.9187 | -35.8474 | -35.8052
25.9248 | -35.8477 | -35.8106
25.9309 | -35.8481 | -35.816
25.937 | -35.8485 | -35.8216
25.9431 | -35.8489 | -35.827
25.9492 | -35.8493 | -35.8323
25.9553 | -35.8496 | -35.8376
25.9614 | -35.85 | -35.8428
25.9675 | -35.8504 | -35.848
25.9736 | -35.8508 | -35.8534
25.9797 | -35.8512 | -35.859
25.9858 | -35.8516 | -35.8644
25.9919 | -35.852 | -35.8699
25.998 | -35.8523 | -35.8752
26.0042 | -35.8527 | -35.8806
26.0103 | -35.8531 | -35.8859
26.0164 | -35.8534 | -35.8913
26.0225 | -35.8538 | -35.8967
26.0286 | -35.8542 | -35.9022
26.0347 | -35.8545 | -35.9074
26.0408 | -35.8549 | -35.9129
26.0469 | -35.8553 | -35.9182
26.053 | -35.8557 | -35.9236
26.0591 | -35.8561 | -35.9292
26.0652 | -35.8564 | -35.9343
26.0713 | -35.8568 | -35.9397
26.0774 | -35.8572 | -35.9452
26.0835 | -35.8575 | -35.9504
26.0896 | -35.8579 | -35.9558
26.0957 | -35.8583 | -35.9611
26.1018 | -35.8586 | -35.9663
26.1079 | -35.859 | -35.9715
26.114 | -35.8594 | -35.977
26.1201 | -35.8597 | -35.9824
26.1262 | -35.8601 | -35.988
26.1323 | -35.8605 | -35.9932
26.1384 | -35.8609 | -35.9985
26.1445 | -35.8612 | -36.004
26.1507 | -35.8616 | -36.0092
26.1568 | -35.862 | -36.0146
26.1629 | -35.8623 | -36.0199
26.169 | -35.8627 | -36.0252
26.1751 | -35.8631 | -36.0303
26.1812 | -35.8635 | -36.0356
26.1873 | -35.8638 | -36.041
26.1934 | -35.8642 | -36.0462
26.1995 | -35.8646 | -36.0515
26.2056 | -35.865 | -36.0567
26.2117 | -35.8653 | -36.0621
26.2178 | -35.8657 | -36.0674
26.2239 | -35.8661 | -36.0726
26.23 | -35.8665 | -36.0782
26.2361 | -35.8669 | -36.0837
26.2422 | -35.8672 | -36.089
26.2483 | -35.8676 | -36.0942
26.2544 | -35.868 | -36.0993
26.2605 | -35.8684 | -36.1048
26.2666 | -35.8687 | -36.1102
26.2727 | -35.8691 | -36.1155
26.2788 | -35.8695 | -36.1206
26.2849 | -35.8698 | -36.1259
26.2911 | -35.8702 | -36.1311
26.2972 | -35.8706 | -36.1365
26.3033 | -35.8709 | -36.1417
26.3094 | -35.8713 | -36.1471
26.3155 | -35.8717 | -36.1521
26.3216 | -35.8721 | -36.1576
26.3277 | -35.8724 | -36.1627
26.3338 | -35.8728 | -36.1678
26.3399 | -35.8731 | -36.1729
26.346 | -35.8735 | -36.178
26.3521 | -35.8739 | -36.1835
26.3582 | -35.8742 | -36.1888
26.3643 | -35.8746 | -36.1941
26.3704 | -35.875 | -36.1995
26.3765 | -35.8754 | -36.205
26.3826 | -35.8758 | -36.2103
26.3887 | -35.8762 | -36.2157
26.3948 | -35.8765 | -36.221
26.4009 | -35.8769 | -36.2264
26.407 | -35.8772 | -36.2316
26.4131 | -35.8776 | -36.2368
26.4192 | -35.878 | -36.242
26.4253 | -35.8784 | -36.2474
26.4314 | -35.8787 | -36.2525
26.4376 | -35.8791 | -36.2577
26.4437 | -35.8795 | -36.2631
26.4498 | -35.8798 | -36.2684
26.4559 | -35.8802 | -36.2735
26.462 | -35.8805 | -36.2788
26.4681 | -35.8809 | -36.2841
26.4742 | -35.8813 | -36.2897
26.4803 | -35.8817 | -36.2949
26.4864 | -35.8821 | -36.3002
26.4925 | -35.8824 | -36.3055
26.4986 | -35.8828 | -36.3109
26.5047 | -35.8832 | -36.3164
26.5108 | -35.8836 | -36.3217
26.5169 | -35.8839 | -36.3269
26.523 | -35.8843 | -36.3323
26.5291 | -35.8847 | -36.3375
26.5352 | -35.885 | -36.343
26.5413 | -35.8854 | -36.3485
26.5474 | -35.8857 | -36.3539
26.5535 | -35.8861 | -36.3593
26.5596 | -35.8864 | -36.3644
26.5657 | -35.8868 | -36.3696
26.5718 | -35.8872 | -36.3749
26.578 | -35.8875 | -36.3802
26.5841 | -35.8879 | -36.3856
26.5902 | -35.8882 | -36.3908
26.5963 | -35.8886 | -36.3959
26.6024 | -35.889 | -36.4013
26.6085 | -35.8894 | -36.4065
26.6146 | -35.8898 | -36.4116
26.6207 | -35.8901 | -36.4169
26.6268 | -35.8905 | -36.4223
26.6329 | -35.8908 | -36.4276
26.639 | -35.8912 | -36.4331
26.6451 | -35.8915 | -36.4385
26.6512 | -35.8919 | -36.4438
26.6573 | -35.8922 | -36.4491
26.6634 | -35.8926 | -36.4542
26.6695 | -35.8929 | -36.4594
26.6756 | -35.8933 | -36.4648
26.6817 | -35.8936 | -36.4701
26.6878 | -35.894 | -36.4755
26.6939 | -35.8944 | -36.4807
26.7 | -35.8947 | -36.4859
26.7061 | -35.8951 | -36.4911
26.7122 | -35.8955 | -36.4964
26.7183 | -35.8958 | -36.5018
26.7245 | -35.8962 | -36.5071
26.7306 | -35.8966 | -36.5124
26.7367 | -35.897 | -36.5175
26.7428 | -35.8974 | -36.5227
26.7489 | -35.8977 | -36.5279
26.755 | -35.8981 | -36.5331
26.7611 | -35.8985 | -36.5384
26.7672 | -35.8988 | -36.5437
26.7733 | -35.8992 | -36.5489
26.7794 | -35.8996 | -36.5544
26.7855 | -35.9 | -36.5595
26.7916 | -35.9003 | -36.5646
26.7977 | -35.9007 | -36.57
26.8038 | -35.9011 | -36.575
26.8099 | -35.9015 | -36.5803
26.816 | -35.9018 | -36.5856
26.8221 | -35.9022 | -36.5908
26.8282 | -35.9025 | -36.596
26.8343 | -35.9029 | -36.6013
26.8404 | -35.9033 | -36.6064
26.8465 | -35.9036 | -36.6117
26.8526 | -35.904 | -36.6168
26.8587 | -35.9044 | -36.6221
26.8649 | -35.9047 | -36.6273
26.871 | -35.9051 | -36.6327
26.8771 | -35.9055 | -36.6379
26.8832 | -35.9058 | -36.643
26.8893 | -35.9062 | -36.6483
26.8954 | -35.9066 | -36.6536
26.9015 | -35.9069 | -36.659
26.9076 | -35.9073 | -36.6644
26.9137 | -35.9077 | -36.6698
26.9198 | -35.9081 | -36.6752
26.9259 | -35.9084 | -36.6804
26.932 | -35.9088 | -36.6855
26.9381 | -35.9092 | -36.691
26.9442 | -35.9094 | -36.696
26.9503 | -35.9098 | -36.7013
26.9564 | -35.9102 | -36.7066
26.9625 | -35.9106 | -36.7118
26.9686 | -35.911 | -36.7171
26.9747 | -35.9114 | -36.7225
26.9808 | -35.9117 | -36.7276
26.9869 | -35.912 | -36.7325
26.993 | -35.9124 | -36.7376
26.9991 | -35.9128 | -36.7427
27.0052 | -35.9131 | -36.7482
27.0114 | -35.9135 | -36.7536
27.0175 | -35.9138 | -36.7587
27.0236 | -35.9142 | -36.7639
27.0297 | -35.9145 | -36.7692
27.0358 | -35.9149 | -36.7745
27.0419 | -35.9153 | -36.7798
27.048 | -35.9156 | -36.7848
27.0541 | -35.916 | -36.7901
27.0602 | -35.9163 | -36.7953
27.0663 | -35.9167 | -36.8006
27.0724 | -35.9171 | -36.8058
27.0785 | -35.9174 | -36.8111
27.0846 | -35.9178 | -36.8161
27.0907 | -35.9181 | -36.8214
27.0968 | -35.9185 | -36.8267
27.1029 | -35.9189 | -36.8317
27.109 | -35.9193 | -36.8368
27.1151 | -35.9196 | -36.8418
27.1212 | -35.92 | -36.8465
27.1273 | -35.9203 | -36.8518
27.1334 | -35.9207 | -36.8572
27.1395 | -35.9211 | -36.8623
27.1456 | -35.9214 | -36.8677
27.1518 | -35.9218 | -36.8728
27.1579 | -35.9221 | -36.8779
27.164 | -35.9225 | -36.8832
27.1701 | -35.9228 | -36.8885
27.1762 | -35.9231 | -36.8936
27.1823 | -35.9235 | -36.8988
27.1884 | -35.9238 | -36.904
27.1945 | -35.9242 | -36.9091
27.2006 | -35.9246 | -36.9141
27.2067 | -35.9249 | -36.9191
27.2128 | -35.9253 | -36.9244
27.2189 | -35.9256 | -36.9297
27.225 | -35.926 | -36.9349
27.2311 | -35.9263 | -36.9399
27.2372 | -35.9267 | -36.9452
27.2433 | -35.9271 | -36.9503
27.2494 | -35.9275 | -36.9556
27.2555 | -35.9278 | -36.9609
27.2616 | -35.9282 | -36.966
27.2677 | -35.9286 | -36.9711
27.2738 | -35.9289 | -36.9763
27.2799 | -35.9293 | -36.9817
27.286 | -35.9296 | -36.9867
27.2921 | -35.93 | -36.992
27.2983 | -35.9303 | -36.997
27.3044 | -35.9307 | -37.0022
27.3105 | -35.931 | -37.0074
27.3166 | -35.9314 | -37.0125
27.3227 | -35.9318 | -37.0174
27.3288 | -35.9321 | -37.0223
27.3349 | -35.9324 | -37.0274
27.341 | -35.9328 | -37.0325
27.3471 | -35.9332 | -37.0376
27.3532 | -35.9335 | -37.0425
27.3593 | -35.9339 | -37.0477
27.3654 | -35.9343 | -37.0529
27.3715 | -35.9346 | -37.0578
27.3776 | -35.935 | -37.0632
27.3837 | -35.9354 | -37.0682
27.3898 | -35.9357 | -37.0731
27.3959 | -35.9361 | -37.0781
27.402 | -35.9364 | -37.083
27.4081 | -35.9368 | -37.0881
27.4142 | -35.9372 | -37.0933
27.4203 | -35.9376 | -37.0984
27.4264 | -35.9379 | -37.1034
27.4325 | -35.9383 | -37.1085
27.4387 | -35.9386 | -37.1135
27.4448 | -35.939 | -37.1187
27.4509 | -35.9393 | -37.1241
27.457 | -35.9397 | -37.1292
27.4631 | -35.9401 | -37.1344
27.4692 | -35.9404 | -37.1393
27.4753 | -35.9407 | -37.144
27.4814 | -35.9411 | -37.1491
27.4875 | -35.9414 | -37.1544
27.4936 | -35.9418 | -37.1595
27.4997 | -35.9422 | -37.1646
27.5058 | -35.9425 | -37.1698
27.5119 | -35.9429 | -37.1747
27.518 | -35.9432 | -37.1798
27.5241 | -35.9435 | -37.185
27.5302 | -35.9439 | -37.1903
27.5363 | -35.9443 | -37.1957
27.5424 | -35.9446 | -37.2006
27.5485 | -35.945 | -37.2058
27.5546 | -35.9453 | -37.2107
27.5607 | -35.9457 | -37.2157
27.5668 | -35.9461 | -37.2208
27.5729 | -35.9464 | -37.2259
27.5791 | -35.9468 | -37.2306
27.5852 | -35.9472 | -37.2356
27.5913 | -35.9475 | -37.2409
27.5974 | -35.9479 | -37.2461
27.6035 | -35.9482 | -37.2513
27.6096 | -35.9485 | -37.2562
27.6157 | -35.9489 | -37.2611
27.6218 | -35.9493 | -37.2662
27.6279 | -35.9496 | -37.2716
27.634 | -35.95 | -37.2769
27.6401 | -35.9503 | -37.2819
27.6462 | -35.9507 | -37.2869
27.6523 | -35.9511 | -37.292
27.6584 | -35.9514 | -37.297
27.6645 | -35.9518 | -37.3021
27.6706 | -35.9521 | -37.3072
27.6767 | -35.9524 | -37.3123
27.6828 | -35.9527 | -37.3171
27.6889 | -35.9531 | -37.3224
27.695 | -35.9534 | -37.3277
27.7011 | -35.9538 | -37.3329
27.7072 | -35.9541 | -37.3381
27.7133 | -35.9545 | -37.3434
27.7194 | -35.9548 | -37.3481
27.7256 | -35.9551 | -37.3532
27.7317 | -35.9554 | -37.3582
27.7378 | -35.9558 | -37.3632
27.7439 | -35.9562 | -37.3687
27.75 | -35.9566 | -37.3738
27.7561 | -35.9569 | -37.3787
27.7622 | -35.9573 | -37.3838
27.7683 | -35.9576 | -37.389
27.7744 | -35.958 | -37.3941
27.7805 | -35.9583 | -37.3992
27.7866 | -35.9586 | -37.4043
27.7927 | -35.959 | -37.4093
27.7988 | -35.9593 | -37.4144
27.8049 | -35.9596 | -37.4195
27.811 | -35.96 | -37.4245
27.8171 | -35.9604 | -37.4296
27.8232 | -35.9607 | -37.4342
27.8293 | -35.9611 | -37.4393
27.8354 | -35.9615 | -37.4444
27.8415 | -35.9619 | -37.4495
27.8476 | -35.9622 | -37.4546
27.8537 | -35.9626 | -37.4598
27.8598 | -35.9629 | -37.4647
27.866 | -35.9633 | -37.4697
27.8721 | -35.9636 | -37.4747
27.8782 | -35.964 | -37.4799
27.8843 | -35.9643 | -37.4849
27.8904 | -35.9647 | -37.49
27.8965 | -35.9651 | -37.4953
27.9026 | -35.9654 | -37.5003
27.9087 | -35.9658 | -37.5055
27.9148 | -35.9661 | -37.5102
27.9209 | -35.9664 | -37.5157
27.927 | -35.9668 | -37.5207
27.9331 | -35.9671 | -37.5259
27.9392 | -35.9675 | -37.531
27.9453 | -35.9678 | -37.536
27.9514 | -35.9681 | -37.5411
27.9575 | -35.9685 | -37.5462
27.9636 | -35.9688 | -37.5514
27.9697 | -35.9692 | -37.5562
27.9758 | -35.9695 | -37.5613
27.9819 | -35.9698 | -37.5664
27.988 | -35.9702 | -37.5711
27.9941 | -35.9705 | -37.5763
28.0002 | -35.9709 | -37.581
28.0063 | -35.9712 | -37.5862
28.0125 | -35.9715 | -37.5913
28.0186 | -35.9718 | -37.5962
28.0247 | -35.9722 | -37.6011
28.0308 | -35.9726 | -37.6059
28.0369 | -35.9729 | -37.6105
28.043 | -35.9732 | -37.6159
28.0491 | -35.9736 | -37.6211
28.0552 | -35.9739 | -37.6262
28.0613 | -35.9742 | -37.6313
28.0674 | -35.9746 | -37.6361
28.0735 | -35.9749 | -37.6408

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.635 | -41.7888
21.4908 | -35.6338 | -34.6425
16.4948 | -34.9919 | -27.413
12.6602 | -34.6241 | -21.2288
9.7171 | -34.404 | -16.3794
7.4582 | -34.242 | -12.4404
5.7244 | -34.1238 | -9.2388
4.3936 | -34.0414 | -6.6613
3.3722 | -34.0335 | -4.6284
2.5883 | -34.0022 | -3.0854
1.9866 | -34.0194 | -1.8475
1.5248 | -34.0092 | -0.85797
1.1703 | -34.0337 | -0.21417
1 | -34.0425 | 0.1042
0.89822 | -34.0217 | 0.37753
0.68941 | -33.9892 | 0.96
0.52919 | -33.9852 | 1.4066
0.40612 | -33.9765 | 2.0137
0.31171 | -33.9807 | 2.4715
0.23928 | -33.9852 | 3.2021
0.18358 | -33.9811 | 3.6198
0.14091 | -33.9582 | 4.5163
0.10819 | -33.9871 | 5.684
0.083017 | -34.0191 | 7.1603
0.063705 | -34.0812 | 9.0434
0.048923 | -34.143 | 11.5463
0.037479 | -34.2288 | 14.738
0.028801 | -34.4078 | 18.7584
0.022078 | -34.7498 | 24.0658
0.017023 | -35.2353 | 30.4364
0.013018 | -36.0173 | 38.1424
0.010014 | -37.1231 | 46.0438

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003751 | -9.8524
0.007502 | -26.1234
0.011252 | -34.8921
0.015003 | -39.9982
0.018754 | -44.234
0.022505 | -46.5656
0.026255 | -47.7004
0.030006 | -49.5093
0.033757 | -51.4793
0.037508 | -52.2974
0.041258 | -53.5657
0.045009 | -55.4924
0.04876 | -56.46
0.052511 | -56.5905
0.056261 | -57.4463
0.060012 | -57.416
0.063763 | -58.1946
0.067514 | -59.4267
0.071264 | -60.1018
0.075015 | -61.1004
0.078766 | -60.7551
0.082517 | -60.9411
0.086268 | -61.9627
0.090018 | -62.0695
0.093769 | -62.0682
0.09752 | -62.6962
0.10127 | -64.7546
0.10502 | -66.0926
0.10877 | -64.9041
0.11252 | -64.6832
0.11627 | -65.0423
0.12002 | -65.624
0.12377 | -66.3727
0.12753 | -67.0403
0.13128 | -67.8509
0.13503 | -68.2641
0.13878 | -67.5047
0.14253 | -67.2667
0.14628 | -67.5633
0.15003 | -68.1153
0.15378 | -68.3919
0.15753 | -68.254
0.16128 | -68.5472
0.16503 | -69.0599
0.16878 | -69.6251
0.17253 | -69.5591
0.17629 | -69.7153
0.18004 | -70.4064
0.18379 | -69.4215
0.18754 | -69.6681
0.19129 | -70.5541
0.19504 | -71.874
0.19879 | -71.7623
0.20254 | -70.8806
0.20629 | -70.6982
0.21004 | -70.7419
0.21379 | -70.9696
0.21754 | -71.6617
0.22129 | -72.4439
0.22505 | -73.0149
0.2288 | -72.7795
0.23255 | -72.8111
0.2363 | -72.7102
0.24005 | -72.1279
0.2438 | -73.0304
0.24755 | -72.7116
0.2513 | -72.9756
0.25505 | -73.2557
0.2588 | -74.2228
0.26255 | -75.0286
0.2663 | -74.4725
0.27005 | -73.6478
0.27381 | -74.0361
0.27756 | -74.9704
0.28131 | -75.0792
0.28506 | -74.593
0.28881 | -74.5617
0.29256 | -74.9016
0.29631 | -74.9539
0.30006 | -75.3712
0.30381 | -75.9206
0.30756 | -77.1675
0.31131 | -77.0834
0.31506 | -76.3272
0.31882 | -75.6901
0.32257 | -76.0799
0.32632 | -76.4696
0.33007 | -76.3154
0.33382 | -76.3092
0.33757 | -76.3538
0.34132 | -77.1467
0.34507 | -77.5681
0.34882 | -77.6176
0.35257 | -77.5005
0.35632 | -77.3977
0.36007 | -77.9833
0.36382 | -76.4193
0.36757 | -75.5983
0.37133 | -76.8369
0.37508 | -78.6172
0.37883 | -78.9464
0.38258 | -77.5845
0.38633 | -78.3168
0.39008 | -77.9151
0.39383 | -77.6849
0.39758 | -77.779
0.40133 | -78.0908
0.40508 | -79.0508
0.40883 | -78.1034
0.41258 | -77.6295
0.41634 | -78.8894
0.42008 | -80.0629
0.42384 | -79.0695
0.42759 | -79.2463
0.43134 | -79.0128
0.43509 | -78.9878
0.43884 | -79.1623
0.44259 | -78.5587
0.44634 | -78.4535
0.45009 | -80.147
0.45384 | -79.918
0.45759 | -79.4382
0.46134 | -79.113
0.46509 | -81.0718
0.46885 | -81.1738
0.4726 | -80.5433
0.47635 | -80.3299
0.4801 | -80.2564
0.48385 | -80.5052
0.4876 | -80.1326
0.49135 | -78.5305
0.4951 | -79.0746
0.49885 | -80.9882
0.5026 | -82.1556
0.50635 | -81.6799
0.5101 | -81.568
0.51385 | -81.5242
0.5176 | -82.3128
0.52136 | -83.2777
0.52511 | -82.3506
0.52886 | -81.2789
0.53261 | -80.3921
0.53636 | -81.3177
0.54011 | -81.1681
0.54386 | -80.9955
0.54761 | -80.851
0.55136 | -81.1389
0.55511 | -81.6573
0.55886 | -81.7855
0.56261 | -81.785
0.56637 | -81.6721
0.57012 | -81.557
0.57387 | -81.3788
0.57762 | -81.2925
0.58137 | -81.3938
0.58512 | -81.8847
0.58887 | -82.3327
0.59262 | -82.4752
0.59637 | -82.4279
0.60012 | -82.2526
0.60387 | -82.0517
0.60762 | -82.0179
0.61137 | -81.9439
0.61513 | -82.269
0.61888 | -82.87
0.62263 | -83.2344
0.62638 | -83.386
0.63013 | -83.4273
0.63388 | -83.4971
0.63763 | -83.495
0.64138 | -83.3537
0.64513 | -83.171
0.64888 | -83.1591
0.65263 | -82.8086
0.65638 | -82.6469
0.66013 | -83.3428
0.66388 | -83.7195
0.66764 | -83.3109
0.67139 | -83.3525
0.67514 | -83.9563
0.67889 | -84.0845
0.68264 | -83.5924
0.68639 | -83.1111
0.69014 | -83.3447
0.69389 | -83.9806
0.69764 | -84.272
0.70139 | -84.0317
0.70514 | -83.879
0.70889 | -84.1525
0.71264 | -84.3547
0.7164 | -83.9982
0.72015 | -83.7243
0.7239 | -83.804
0.72765 | -83.9834
0.7314 | -84.0701
0.73515 | -84.0679
0.7389 | -83.6692
0.74265 | -83.6279
0.7464 | -84.2542
0.75015 | -84.7165
0.7539 | -84.7228
0.75765 | -84.749
0.7614 | -84.9068
0.76516 | -84.6355
0.76891 | -84.2003
0.77266 | -84.2929
0.77641 | -85.0248
0.78016 | -86.0928
0.78391 | -86.6812
0.78766 | -86.4162
0.79141 | -86.0039
0.79516 | -85.8063
0.79891 | -85.5293
0.80266 | -85.13
0.80641 | -85.1349
0.81017 | -85.2799
0.81392 | -85.2331
0.81767 | -85.3201
0.82142 | -85.6759
0.82517 | -86.1654
0.82892 | -86.3363
0.83267 | -86.0968
0.83642 | -85.7542
0.84017 | -85.5009
0.84392 | -85.448
0.84767 | -85.6607
0.85142 | -85.6106
0.85517 | -85.362
0.85892 | -85.3551
0.86267 | -85.8892
0.86643 | -86.327
0.87018 | -86.5742
0.87393 | -86.6996
0.87768 | -86.0793
0.88143 | -85.5391
0.88518 | -85.48
0.88893 | -85.834
0.89268 | -86.4208
0.89643 | -86.6779
0.90018 | -86.2473
0.90393 | -86.1437
0.90768 | -86.5171
0.91143 | -86.8537
0.91519 | -86.7164
0.91894 | -86.5694
0.92269 | -86.7335
0.92644 | -87.1581
0.93019 | -87.2638
0.93394 | -87.1574
0.93769 | -87.3225
0.94144 | -87.5134
0.94519 | -87.145
0.94894 | -86.6613
0.95269 | -86.8076
0.95644 | -87.1799
0.9602 | -87.7702
0.96395 | -88.3463
0.9677 | -88.3147
0.97145 | -88.1218
0.9752 | -87.9472
0.97895 | -87.5347
0.9827 | -87.3115
0.98645 | -87.3916
0.9902 | -87.9856
0.99395 | -88.6319
0.9977 | -88.323
1.0015 | -87.5762
1.0052 | -87.1839
1.009 | -87.2542
1.0127 | -87.3171
1.0165 | -87.4925
1.0202 | -87.7465
1.024 | -87.8158
1.0277 | -87.9325
1.0315 | -88.2202
1.0352 | -88.1861
1.039 | -88.0874
1.0427 | -88.2448
1.0465 | -88.4164
1.0502 | -88.5547
1.054 | -88.5808
1.0577 | -88.6599
1.0615 | -89.0854
1.0652 | -89.0425
1.069 | -88.4616
1.0727 | -88.0529
1.0765 | -87.9071
1.0802 | -88.1765
1.084 | -88.4423
1.0877 | -88.5741
1.0915 | -88.3972
1.0952 | -88.1166
1.099 | -88.1407
1.1027 | -88.4551
1.1065 | -88.6373
1.1102 | -88.5115
1.114 | -88.3237
1.1177 | -88.3377
1.1215 | -88.5605
1.1252 | -88.7911
1.129 | -88.843
1.1327 | -88.7177
1.1365 | -88.6138
1.1402 | -88.6645
1.144 | -88.8698
1.1477 | -89.1379
1.1515 | -89.3153
1.1552 | -89.3082
1.159 | -89.2023
1.1627 | -89.1926
1.1665 | -89.3089
1.1702 | -89.3735
1.174 | -89.2874
1.1777 | -89.1677
1.1815 | -89.1033
1.1852 | -89.0628
1.189 | -89.0809
1.1927 | -89.1924
1.1965 | -89.2493
1.2002 | -89.1667
1.204 | -89.1289
1.2077 | -89.29
1.2115 | -89.5487
1.2152 | -89.6817
1.219 | -89.5678
1.2227 | -89.3332
1.2265 | -89.1999
1.2303 | -89.1635
1.234 | -89.1252
1.2378 | -89.1697
1.2415 | -89.2623
1.2453 | -89.1488
1.249 | -88.9247
1.2528 | -88.8914
1.2565 | -89.1328
1.2603 | -89.5073
1.264 | -89.7433
1.2678 | -89.7403
1.2715 | -89.6011
1.2753 | -89.4198
1.279 | -89.3088
1.2828 | -89.374
1.2865 | -89.6379
1.2903 | -89.9302
1.294 | -90.0816
1.2978 | -90.1129
1.3015 | -90.103
1.3053 | -90.0922
1.309 | -90.0643
1.3128 | -90.0898
1.3165 | -90.2801
1.3203 | -90.6118
1.324 | -90.9224
1.3278 | -90.9621
1.3315 | -90.7587
1.3353 | -90.5524
1.339 | -90.4439
1.3428 | -90.4251
1.3465 | -90.4709
1.3503 | -90.5664
1.354 | -90.773
1.3578 | -91.0677
1.3615 | -91.263
1.3653 | -91.3169
1.369 | -91.3239
1.3728 | -91.1539
1.3765 | -90.7642
1.3803 | -90.4074
1.384 | -90.2621
1.3878 | -90.2384
1.3915 | -90.2365
1.3953 | -90.3625
1.399 | -90.6134
1.4028 | -90.7996
1.4065 | -90.8985
1.4103 | -90.9696
1.414 | -91.013
1.4178 | -91.0517
1.4215 | -91.1188
1.4253 | -91.1656
1.429 | -91.077
1.4328 | -90.808
1.4365 | -90.4466
1.4403 | -90.1686
1.444 | -90.1778
1.4478 | -90.5666
1.4515 | -91.1139
1.4553 | -91.3773
1.459 | -91.338
1.4628 | -91.3296
1.4665 | -91.4114
1.4703 | -91.3982
1.474 | -91.1676
1.4778 | -90.8774
1.4816 | -90.765
1.4853 | -90.8469
1.4891 | -90.9611
1.4928 | -91.0177
1.4966 | -90.9665
1.5003 | -90.8758
1.5041 | -91.01
1.5078 | -91.4483
1.5116 | -91.8792
1.5153 | -92.0048
1.5191 | -91.7607
1.5228 | -91.3623
1.5266 | -91.2356
1.5303 | -91.5349
1.5341 | -91.828
1.5378 | -91.5325
1.5416 | -90.887
1.5453 | -90.5481
1.5491 | -90.7011
1.5528 | -91.0867
1.5566 | -91.3276
1.5603 | -91.2832
1.5641 | -91.1083
1.5678 | -91.0222
1.5716 | -91.1463
1.5753 | -91.3577
1.5791 | -91.4467
1.5828 | -91.4648
1.5866 | -91.6101
1.5903 | -91.9587
1.5941 | -92.3505
1.5978 | -92.5489
1.6016 | -92.5838
1.6053 | -92.5213
1.6091 | -92.2869
1.6128 | -91.9398
1.6166 | -91.694
1.6203 | -91.7911
1.6241 | -91.8821
1.6278 | -91.9907
1.6316 | -91.9853
1.6353 | -91.8442
1.6391 | -91.6699
1.6428 | -91.5865
1.6466 | -91.6036
1.6503 | -91.6836
1.6541 | -91.8367
1.6578 | -92.0481
1.6616 | -92.2234
1.6653 | -92.247
1.6691 | -92.0755
1.6728 | -91.7713
1.6766 | -91.4218
1.6803 | -91.1151
1.6841 | -90.9513
1.6878 | -91.0347
1.6916 | -91.3863
1.6953 | -91.8335
1.6991 | -92.12
1.7028 | -92.1185
1.7066 | -91.9219
1.7103 | -91.7536
1.7141 | -91.7832
1.7178 | -92.0095
1.7216 | -92.2975
1.7253 | -92.4956
1.7291 | -92.5739
1.7329 | -92.6006
1.7366 | -92.6105
1.7404 | -92.5756
1.7441 | -92.4233
1.7479 | -92.2311
1.7516 | -92.1392
1.7554 | -92.2024
1.7591 | -92.3656
1.7629 | -92.5286
1.7666 | -92.6026
1.7704 | -92.5467
1.7741 | -92.3897
1.7779 | -92.2219
1.7816 | -92.1092
1.7854 | -92.1187
1.7891 | -92.256
1.7929 | -92.4433
1.7966 | -92.5562
1.8004 | -92.5457
1.8041 | -92.4727
1.8079 | -92.4124
1.8116 | -92.4514
1.8154 | -92.6107
1.8191 | -92.8033
1.8229 | -92.9213
1.8266 | -92.9178
1.8304 | -92.8087
1.8341 | -92.6474
1.8379 | -92.4913
1.8416 | -92.3855
1.8454 | -92.3557
1.8491 | -92.3838
1.8529 | -92.4234
1.8566 | -92.4378
1.8604 | -92.4482
1.8641 | -92.5012
1.8679 | -92.5729
1.8716 | -92.65
1.8754 | -92.7584
1.8791 | -92.9527
1.8829 | -93.2786
1.8866 | -93.6731
1.8904 | -93.9876
1.8941 | -94.1019
1.8979 | -94.0078
1.9016 | -93.8199
1.9054 | -93.655
1.9091 | -93.5681
1.9129 | -93.4769
1.9166 | -93.3189
1.9204 | -93.1725
1.9241 | -93.1441
1.9279 | -93.2474
1.9316 | -93.4164
1.9354 | -93.6137
1.9391 | -93.8062
1.9429 | -93.9243
1.9466 | -93.904
1.9504 | -93.7546
1.9541 | -93.5821
1.9579 | -93.5147
1.9616 | -93.5967
1.9654 | -93.7734
1.9691 | -93.9311
1.9729 | -93.9936
1.9767 | -93.9454
1.9804 | -93.8242
1.9842 | -93.6779
1.9879 | -93.5618
1.9917 | -93.5142
1.9954 | -93.5504
1.9992 | -93.6568
2.0029 | -93.8023
2.0067 | -93.9448
2.0104 | -94.0182
2.0142 | -93.9718
2.0179 | -93.8201
2.0217 | -93.6354
2.0254 | -93.5059
2.0292 | -93.4759
2.0329 | -93.5195
2.0367 | -93.5507
2.0404 | -93.5068
2.0442 | -93.4297
2.0479 | -93.3996
2.0517 | -93.4294
2.0554 | -93.4644
2.0592 | -93.5204
2.0629 | -93.6772
2.0667 | -93.926
2.0704 | -94.1713
2.0742 | -94.3642
2.0779 | -94.5511
2.0817 | -94.7747
2.0854 | -94.9551
2.0892 | -94.9378
2.0929 | -94.6704
2.0967 | -94.3032
2.1004 | -94.029
2.1042 | -93.9022
2.1079 | -93.8861
2.1117 | -93.9211
2.1154 | -93.9538
2.1192 | -93.9541
2.1229 | -93.9198
2.1267 | -93.8841
2.1304 | -93.8692
2.1342 | -93.8599
2.1379 | -93.8581
2.1417 | -93.9235
2.1454 | -94.1168
2.1492 | -94.3736
2.1529 | -94.5317
2.1567 | -94.4825
2.1604 | -94.2821
2.1642 | -94.1811
2.1679 | -94.169
2.1717 | -94.2713
2.1754 | -94.446
2.1792 | -94.5929
2.1829 | -94.6083
2.1867 | -94.473
2.1904 | -94.2383
2.1942 | -94.0237
2.1979 | -93.9246
2.2017 | -93.9629
2.2054 | -94.073
2.2092 | -94.1688
2.213 | -94.1698
2.2167 | -94.0859
2.2205 | -93.9495
2.2242 | -93.8105
2.228 | -93.7139
2.2317 | -93.6927
2.2355 | -93.7445
2.2392 | -93.8456
2.243 | -93.9811
2.2467 | -94.1386
2.2505 | -94.3235
2.2542 | -94.5307
2.258 | -94.7174
2.2617 | -94.8446
2.2655 | -94.8756
2.2692 | -94.8219
2.273 | -94.7272
2.2767 | -94.6604
2.2805 | -94.6764
2.2842 | -94.7472
2.288 | -94.8221
2.2917 | -94.8607
2.2955 | -94.8653
2.2992 | -94.8561
2.303 | -94.8492
2.3067 | -94.8428
2.3105 | -94.8354
2.3142 | -94.8487
2.318 | -94.8829
2.3217 | -94.9471
2.3255 | -95.0194
2.3292 | -95.04
2.333 | -95.0201
2.3367 | -95.0172
2.3405 | -95.0589
2.3442 | -95.1082
2.348 | -95.1007
2.3517 | -95.017
2.3555 | -94.9049
2.3592 | -94.8281
2.363 | -94.7947
2.3667 | -94.7962
2.3705 | -94.834
2.3742 | -94.9082
2.378 | -95.0092
2.3817 | -95.0708
2.3855 | -95.0054
2.3892 | -94.817
2.393 | -94.5971
2.3967 | -94.4379
2.4005 | -94.3838
2.4042 | -94.4338
2.408 | -94.5531
2.4117 | -94.6977
2.4155 | -94.828
2.4192 | -94.932
2.423 | -95.0222
2.4267 | -95.1134
2.4305 | -95.1741
2.4342 | -95.1651
2.438 | -95.0566
2.4417 | -94.8862
2.4455 | -94.7249
2.4492 | -94.6426
2.453 | -94.6499
2.4567 | -94.7182
2.4605 | -94.8098
2.4643 | -94.9
2.468 | -94.9893
2.4718 | -95.0903
2.4755 | -95.218
2.4793 | -95.3819
2.483 | -95.5432
2.4868 | -95.6659
2.4905 | -95.7187
2.4943 | -95.7034
2.498 | -95.6325
2.5018 | -95.4813
2.5055 | -95.247
2.5093 | -94.9546
2.513 | -94.6644
2.5168 | -94.4514
2.5205 | -94.3801
2.5243 | -94.4839
2.528 | -94.738
2.5318 | -95.0733
2.5355 | -95.3844
2.5393 | -95.5803
2.543 | -95.6566
2.5468 | -95.6415
2.5505 | -95.6052
2.5543 | -95.6057
2.558 | -95.6641
2.5618 | -95.7314
2.5655 | -95.7227
2.5693 | -95.59
2.573 | -95.3756
2.5768 | -95.1714
2.5805 | -95.0447
2.5843 | -95.0165
2.588 | -95.0569
2.5918 | -95.1267
2.5955 | -95.2012
2.5993 | -95.2675
2.603 | -95.3246
2.6068 | -95.343
2.6105 | -95.2938
2.6143 | -95.2042
2.618 | -95.1423
2.6218 | -95.1707
2.6255 | -95.2985
2.6293 | -95.5053
2.633 | -95.7187
2.6368 | -95.8433
2.6405 | -95.8234
2.6443 | -95.6743
2.648 | -95.4756
2.6518 | -95.3216
2.6555 | -95.2456
2.6593 | -95.2434
2.663 | -95.2818
2.6668 | -95.3652
2.6705 | -95.5033
2.6743 | -95.6883
2.678 | -95.858
2.6818 | -95.9003
2.6855 | -95.7913
2.6893 | -95.6249
2.693 | -95.5104
2.6968 | -95.524
2.7005 | -95.6741
2.7043 | -95.8978
2.7081 | -96.0461
2.7118 | -96.128
2.7156 | -96.1031
2.7193 | -95.99
2.7231 | -95.841
2.7268 | -95.7195
2.7306 | -95.6692
2.7343 | -95.6945
2.7381 | -95.7735
2.7418 | -95.8682
2.7456 | -95.9318
2.7493 | -95.9411
2.7531 | -95.8862
2.7568 | -95.7878
2.7606 | -95.6936
2.7643 | -95.6409
2.7681 | -95.6458
2.7718 | -95.7011
2.7756 | -95.7792
2.7793 | -95.8359
2.7831 | -95.8582
2.7868 | -95.8666
2.7906 | -95.8695
2.7943 | -95.8817
2.7981 | -95.9149
2.8018 | -95.9651
2.8056 | -96.0266
2.8093 | -96.1026
2.8131 | -96.1842
2.8168 | -96.2595
2.8206 | -96.3242
2.8243 | -96.3581
2.8281 | -96.3582
2.8318 | -96.3163
2.8356 | -96.2472
2.8393 | -96.1857
2.8431 | -96.1511
2.8468 | -96.1438
2.8506 | -96.1805
2.8543 | -96.2543
2.8581 | -96.3702
2.8618 | -96.5222
2.8656 | -96.6773
2.8693 | -96.8007
2.8731 | -96.8626
2.8768 | -96.8624
2.8806 | -96.8244
2.8843 | -96.7775
2.8881 | -96.7382
2.8918 | -96.7157
2.8956 | -96.7216
2.8993 | -96.7586
2.9031 | -96.8222
2.9068 | -96.8941
2.9106 | -96.9526
2.9143 | -96.9666
2.9181 | -96.9214
2.9218 | -96.8362
2.9256 | -96.741
2.9293 | -96.6683
2.9331 | -96.6182
2.9368 | -96.5777
2.9406 | -96.5365
2.9443 | -96.4897
2.9481 | -96.4606
2.9518 | -96.465
2.9556 | -96.4905
2.9594 | -96.5361
2.9631 | -96.6009
2.9669 | -96.6588
2.9706 | -96.702
2.9744 | -96.7237
2.9781 | -96.7248
2.9819 | -96.7231
2.9856 | -96.7079
2.9894 | -96.6662
2.9931 | -96.6002
2.9969 | -96.5238
3.0006 | -96.464
3.0044 | -96.4471
3.0081 | -96.47
3.0119 | -96.5515
3.0156 | -96.6921
3.0194 | -96.8373
3.0231 | -96.9462
3.0269 | -96.9723
3.0306 | -96.9132
3.0344 | -96.8215
3.0381 | -96.7444
3.0419 | -96.7335
3.0456 | -96.8081
3.0494 | -96.9775
3.0531 | -97.2049
3.0569 | -97.4535
3.0606 | -97.6478
3.0644 | -97.7401
3.0681 | -97.7227
3.0719 | -97.5977
3.0756 | -97.4217
3.0794 | -97.2305
3.0831 | -97.0712
3.0869 | -96.9861
3.0906 | -96.9792
3.0944 | -97.0282
3.0981 | -97.086
3.1019 | -97.1044
3.1056 | -97.0524
3.1094 | -96.9636
3.1131 | -96.8883
3.1169 | -96.8701
3.1206 | -96.9295
3.1244 | -97.0459
3.1281 | -97.1818
3.1319 | -97.2753
3.1356 | -97.3031
3.1394 | -97.2792
3.1431 | -97.2268
3.1469 | -97.1737
3.1506 | -97.1373
3.1544 | -97.1129
3.1581 | -97.0891
3.1619 | -97.0641
3.1656 | -97.0408
3.1694 | -97.0423
3.1731 | -97.1025
3.1769 | -97.2445
3.1806 | -97.4821
3.1844 | -97.7905
3.1881 | -98.1152
3.1919 | -98.4101
3.1956 | -98.579
3.1994 | -98.5777
3.2032 | -98.4187
3.2069 | -98.1507
3.2107 | -97.8372
3.2144 | -97.5333
3.2182 | -97.272
3.2219 | -97.0908
3.2257 | -96.9988
3.2294 | -96.9781
3.2332 | -97.0047
3.2369 | -97.0561
3.2407 | -97.1042
3.2444 | -97.1579
3.2482 | -97.2205
3.2519 | -97.2793
3.2557 | -97.3232
3.2594 | -97.3575
3.2632 | -97.3738
3.2669 | -97.3736
3.2707 | -97.3617
3.2744 | -97.3247
3.2782 | -97.2876
3.2819 | -97.2704
3.2857 | -97.2723
3.2894 | -97.302
3.2932 | -97.3516
3.2969 | -97.4033
3.3007 | -97.4461
3.3044 | -97.4889
3.3082 | -97.545
3.3119 | -97.6129
3.3157 | -97.6893
3.3194 | -97.7759
3.3232 | -97.8461
3.3269 | -97.8991
3.3307 | -97.9332
3.3344 | -97.9418
3.3382 | -97.9335
3.3419 | -97.9189
3.3457 | -97.9068
3.3494 | -97.9078
3.3532 | -97.9237
3.3569 | -97.9517
3.3607 | -97.9684
3.3644 | -97.9753
3.3682 | -97.9563
3.3719 | -97.9072
3.3757 | -97.8254
3.3794 | -97.7123
3.3832 | -97.5751
3.3869 | -97.4528
3.3907 | -97.3642
3.3944 | -97.3127
3.3982 | -97.2958
3.4019 | -97.3107
3.4057 | -97.3562
3.4094 | -97.4284
3.4132 | -97.5245
3.4169 | -97.6331
3.4207 | -97.7387
3.4244 | -97.8251
3.4282 | -97.8719
3.4319 | -97.8966
3.4357 | -97.9122
3.4394 | -97.901
3.4432 | -97.8598
3.447 | -97.8059
3.4507 | -97.7503
3.4545 | -97.6963
3.4582 | -97.6484
3.462 | -97.6118
3.4657 | -97.6034
3.4695 | -97.6058
3.4732 | -97.6335
3.477 | -97.6745
3.4807 | -97.7208
3.4845 | -97.7528
3.4882 | -97.7464
3.492 | -97.6998
3.4957 | -97.6379
3.4995 | -97.5927
3.5032 | -97.5804
3.507 | -97.6328
3.5107 | -97.758
3.5145 | -97.9354
3.5182 | -98.1328
3.522 | -98.2949
3.5257 | -98.3947
3.5295 | -98.3908
3.5332 | -98.3049
3.537 | -98.1486
3.5407 | -97.9615
3.5445 | -97.7782
3.5482 | -97.6226
3.552 | -97.5029
3.5557 | -97.4175
3.5595 | -97.3698
3.5632 | -97.3392
3.567 | -97.3229
3.5707 | -97.3177
3.5745 | -97.3314
3.5782 | -97.3671
3.582 | -97.4286
3.5857 | -97.525
3.5895 | -97.6592
3.5932 | -97.8431
3.597 | -98.0544
3.6007 | -98.2725
3.6045 | -98.476
3.6082 | -98.6332
3.612 | -98.7277
3.6157 | -98.7628
3.6195 | -98.7568
3.6232 | -98.7167
3.627 | -98.6609
3.6307 | -98.5965
3.6345 | -98.5326
3.6382 | -98.4816
3.642 | -98.4505
3.6457 | -98.4487
3.6495 | -98.4744
3.6532 | -98.5168
3.657 | -98.574
3.6607 | -98.64
3.6645 | -98.7156
3.6682 | -98.7939
3.672 | -98.8597
3.6757 | -98.9019
3.6795 | -98.921
3.6832 | -98.9243
3.687 | -98.9373
3.6907 | -98.9868
3.6945 | -99.0738
3.6983 | -99.1722
3.702 | -99.2539
3.7058 | -99.2612
3.7095 | -99.1945
3.7133 | -99.048
3.717 | -98.8564
3.7208 | -98.6523
3.7245 | -98.4702
3.7283 | -98.3371
3.732 | -98.2557
3.7358 | -98.2196
3.7395 | -98.2201
3.7433 | -98.2522
3.747 | -98.3057
3.7508 | -98.377
3.7545 | -98.4562
3.7583 | -98.5313
3.762 | -98.6042
3.7658 | -98.6583
3.7695 | -98.7005
3.7733 | -98.7355
3.777 | -98.7745
3.7808 | -98.8122
3.7845 | -98.8556
3.7883 | -98.8991
3.792 | -98.9479
3.7958 | -99.0046
3.7995 | -99.0546
3.8033 | -99.0928
3.807 | -99.114
3.8108 | -99.1157
3.8145 | -99.0956
3.8183 | -99.0458
3.822 | -98.9829
3.8258 | -98.9046
3.8295 | -98.8261
3.8333 | -98.7732
3.837 | -98.7505
3.8408 | -98.7737
3.8445 | -98.8408
3.8483 | -98.9371
3.852 | -99.0498
3.8558 | -99.1656
3.8595 | -99.2572
3.8633 | -99.3231
3.867 | -99.3478
3.8708 | -99.3334
3.8745 | -99.2877
3.8783 | -99.2268
3.882 | -99.1717
3.8858 | -99.1387
3.8895 | -99.1346
3.8933 | -99.1469
3.897 | -99.1583
3.9008 | -99.1739
3.9045 | -99.1686
3.9083 | -99.1418
3.912 | -99.0927
3.9158 | -99.044
3.9195 | -99.0093
3.9233 | -98.9995
3.927 | -99.004
3.9308 | -99.0178
3.9345 | -99.0428
3.9383 | -99.0635
3.9421 | -99.081
3.9458 | -99.0967
3.9496 | -99.1273
3.9533 | -99.1786
3.9571 | -99.243
3.9608 | -99.3114
3.9646 | -99.3662
3.9683 | -99.4142
3.9721 | -99.4495
3.9758 | -99.4625
3.9796 | -99.4517
3.9833 | -99.424
3.9871 | -99.3882
3.9908 | -99.3534
3.9946 | -99.3155
3.9983 | -99.2923
4.0021 | -99.303
4.0058 | -99.3355
4.0096 | -99.3715
4.0133 | -99.4011
4.0171 | -99.4252
4.0208 | -99.443
4.0246 | -99.4538
4.0283 | -99.4712
4.0321 | -99.4976
4.0358 | -99.5288
4.0396 | -99.5547
4.0433 | -99.5653
4.0471 | -99.549
4.0508 | -99.5244
4.0546 | -99.4937
4.0583 | -99.4633
4.0621 | -99.4138
4.0658 | -99.3625
4.0696 | -99.3256
4.0733 | -99.3037
4.0771 | -99.2983
4.0808 | -99.3061
4.0846 | -99.317
4.0883 | -99.3146
4.0921 | -99.292
4.0958 | -99.2506
4.0996 | -99.2003
4.1033 | -99.1427
4.1071 | -99.0774
4.1108 | -99.0212
4.1146 | -98.9637
4.1183 | -98.9343
4.1221 | -98.9453
4.1258 | -98.9909
4.1296 | -99.0665
4.1333 | -99.1579
4.1371 | -99.2611
4.1408 | -99.3581
4.1446 | -99.4453
4.1483 | -99.5285
4.1521 | -99.5978
4.1558 | -99.6481
4.1596 | -99.6776
4.1633 | -99.6752
4.1671 | -99.6617
4.1708 | -99.6439
4.1746 | -99.6228
4.1783 | -99.6279
4.1821 | -99.654
4.1859 | -99.7004
4.1896 | -99.7657
4.1934 | -99.8262
4.1971 | -99.8825
4.2009 | -99.9249
4.2046 | -99.9464
4.2084 | -99.957
4.2121 | -99.9573
4.2159 | -99.942
4.2196 | -99.8984
4.2234 | -99.8243
4.2271 | -99.716
4.2309 | -99.5751
4.2346 | -99.4342
4.2384 | -99.3196
4.2421 | -99.2342
4.2459 | -99.1882
4.2496 | -99.1906
4.2534 | -99.2378
4.2571 | -99.3332
4.2609 | -99.4683
4.2646 | -99.6341
4.2684 | -99.8313
4.2721 | -100.0177
4.2759 | -100.1704
4.2796 | -100.2725
4.2834 | -100.2997
4.2871 | -100.2483
4.2909 | -100.1571
4.2946 | -100.0591
4.2984 | -99.9482
4.3021 | -99.8483
4.3059 | -99.7658
4.3096 | -99.698
4.3134 | -99.6485
4.3171 | -99.6109
4.3209 | -99.5724
4.3246 | -99.5336
4.3284 | -99.4979
4.3321 | -99.4756
4.3359 | -99.4558
4.3396 | -99.4519
4.3434 | -99.4589
4.3471 | -99.4806
4.3509 | -99.5285
4.3546 | -99.5972
4.3584 | -99.6663
4.3621 | -99.7276
4.3659 | -99.7836
4.3696 | -99.8179
4.3734 | -99.8317
4.3771 | -99.825
4.3809 | -99.7991
4.3846 | -99.7777
4.3884 | -99.7673
4.3921 | -99.7735
4.3959 | -99.787
4.3996 | -99.8058
4.4034 | -99.8352
4.4071 | -99.866
4.4109 | -99.8923
4.4146 | -99.914
4.4184 | -99.9287
4.4221 | -99.922
4.4259 | -99.904
4.4296 | -99.8802
4.4334 | -99.8398
4.4372 | -99.8048
4.4409 | -99.7778
4.4447 | -99.7619
4.4484 | -99.7631
4.4522 | -99.7875
4.4559 | -99.8264
4.4597 | -99.8888
4.4634 | -99.9554
4.4672 | -100.0238
4.4709 | -100.0985
4.4747 | -100.1693
4.4784 | -100.2321
4.4822 | -100.2865
4.4859 | -100.3251
4.4897 | -100.3477
4.4934 | -100.3579
4.4972 | -100.3539
4.5009 | -100.3507
4.5047 | -100.3523
4.5084 | -100.3424
4.5122 | -100.3281
4.5159 | -100.3097
4.5197 | -100.2828
4.5234 | -100.2441
4.5272 | -100.2013
4.5309 | -100.1658
4.5347 | -100.1502
4.5384 | -100.1508
4.5422 | -100.1688
4.5459 | -100.2027
4.5497 | -100.248
4.5534 | -100.2892
4.5572 | -100.3187
4.5609 | -100.342
4.5647 | -100.3642
4.5684 | -100.3781
4.5722 | -100.3846
4.5759 | -100.3965
4.5797 | -100.4109
4.5834 | -100.4337
4.5872 | -100.4533
4.5909 | -100.4655
4.5947 | -100.4778
4.5984 | -100.494
4.6022 | -100.5047
4.6059 | -100.511
4.6097 | -100.5075
4.6134 | -100.4835
4.6172 | -100.4447
4.6209 | -100.396
4.6247 | -100.3518
4.6284 | -100.3243
4.6322 | -100.313
4.6359 | -100.3248
4.6397 | -100.3632
4.6434 | -100.4182
4.6472 | -100.4739
4.6509 | -100.5303
4.6547 | -100.5678
4.6584 | -100.5912
4.6622 | -100.5911
4.6659 | -100.561
4.6697 | -100.5097
4.6734 | -100.4538
4.6772 | -100.395
4.681 | -100.346
4.6847 | -100.3048
4.6885 | -100.278
4.6922 | -100.2655
4.696 | -100.2481
4.6997 | -100.2166
4.7035 | -100.1821
4.7072 | -100.1426
4.711 | -100.1121
4.7147 | -100.1008
4.7185 | -100.1067
4.7222 | -100.121
4.726 | -100.1402
4.7297 | -100.1581
4.7335 | -100.1817
4.7372 | -100.1989
4.741 | -100.2197
4.7447 | -100.2403
4.7485 | -100.2673
4.7522 | -100.296
4.756 | -100.3309
4.7597 | -100.3627
4.7635 | -100.3843
4.7672 | -100.4089
4.771 | -100.4299
4.7747 | -100.435
4.7785 | -100.4313
4.7822 | -100.4014
4.786 | -100.3601
4.7897 | -100.3163
4.7935 | -100.2692
4.7972 | -100.2248
4.801 | -100.1885
4.8047 | -100.1519
4.8085 | -100.1301
4.8122 | -100.124
4.816 | -100.1324
4.8197 | -100.1509
4.8235 | -100.1712
4.8272 | -100.1882
4.831 | -100.2086
4.8347 | -100.2381
4.8385 | -100.2799
4.8422 | -100.3407
4.846 | -100.4166
4.8497 | -100.5031
4.8535 | -100.6017
4.8572 | -100.6904
4.861 | -100.7671
4.8647 | -100.8284
4.8685 | -100.8505
4.8722 | -100.8866
4.876 | -100.9201
4.8797 | -100.9578
4.8835 | -100.9908
4.8872 | -101.0174
4.891 | -101.051
4.8947 | -101.0815
4.8985 | -101.1051
4.9022 | -101.1205
4.906 | -101.1217
4.9097 | -101.1177
4.9135 | -101.0997
4.9172 | -101.0639
4.921 | -101.0195
4.9248 | -100.9684
4.9285 | -100.9105
4.9323 | -100.8532
4.936 | -100.8038
4.9398 | -100.7709
4.9435 | -100.7484
4.9473 | -100.735
4.951 | -100.7369
4.9548 | -100.746
4.9585 | -100.7691
4.9623 | -100.8036
4.966 | -100.8411
4.9698 | -100.8759
4.9735 | -100.9122
4.9773 | -100.9412
4.981 | -100.9561
4.9848 | -100.9586
4.9885 | -100.9512
4.9923 | -100.9484
4.996 | -100.9587
4.9998 | -100.9767
5.0035 | -101.0034
5.0073 | -101.0441
5.011 | -101.0984
5.0148 | -101.1647
5.0185 | -101.2344
5.0223 | -101.3048
5.026 | -101.3693
5.0298 | -101.4171
5.0335 | -101.4375
5.0373 | -101.4222
5.041 | -101.3821
5.0448 | -101.3151
5.0485 | -101.2249
5.0523 | -101.1182
5.056 | -101.0016
5.0598 | -100.8928
5.0635 | -100.8007
5.0673 | -100.7301
5.071 | -100.6877
5.0748 | -100.6776
5.0785 | -100.701
5.0823 | -100.7563
5.086 | -100.8332
5.0898 | -100.9177
5.0935 | -101.0048
5.0973 | -101.0864
5.101 | -101.1522
5.1048 | -101.1929
5.1085 | -101.2009
5.1123 | -101.1815
5.116 | -101.1365
5.1198 | -101.0649
5.1235 | -100.9864
5.1273 | -100.9031
5.131 | -100.8136
5.1348 | -100.7354
5.1385 | -100.6615
5.1423 | -100.5883
5.146 | -100.5251
5.1498 | -100.4722
5.1535 | -100.434
5.1573 | -100.4148
5.161 | -100.4187
5.1648 | -100.4555
5.1685 | -100.5206
5.1723 | -100.6062
5.1761 | -100.7077
5.1798 | -100.8177
5.1836 | -100.927
5.1873 | -101.0331
5.1911 | -101.1225
5.1948 | -101.1954
5.1986 | -101.2602
5.2023 | -101.3108
5.2061 | -101.3548
5.2098 | -101.3984
5.2136 | -101.4342
5.2173 | -101.4695
5.2211 | -101.5124
5.2248 | -101.5535
5.2286 | -101.5887
5.2323 | -101.6151
5.2361 | -101.6249
5.2398 | -101.6126
5.2436 | -101.5802
5.2473 | -101.5408
5.2511 | -101.4966
5.2548 | -101.4637
5.2586 | -101.4485
5.2623 | -101.4505
5.2661 | -101.4723
5.2698 | -101.5069
5.2736 | -101.547
5.2773 | -101.5781
5.2811 | -101.6045
5.2848 | -101.6207
5.2886 | -101.6238
5.2923 | -101.626
5.2961 | -101.6271
5.2998 | -101.6311
5.3036 | -101.6393
5.3073 | -101.6567
5.3111 | -101.6859
5.3148 | -101.7206
5.3186 | -101.7579
5.3223 | -101.7873
5.3261 | -101.8086
5.3298 | -101.8215
5.3336 | -101.8182
5.3373 | -101.8019
5.3411 | -101.7778
5.3448 | -101.7499
5.3486 | -101.7247
5.3523 | -101.695
5.3561 | -101.668
5.3598 | -101.6492
5.3636 | -101.6426
5.3673 | -101.6425
5.3711 | -101.6326
5.3748 | -101.6195
5.3786 | -101.6022
5.3823 | -101.5896
5.3861 | -101.5736
5.3898 | -101.5497
5.3936 | -101.5283
5.3973 | -101.5108
5.4011 | -101.5025
5.4048 | -101.4997
5.4086 | -101.5092
5.4123 | -101.5081
5.4161 | -101.5022
5.4199 | -101.4975
5.4236 | -101.4923
5.4274 | -101.4871
5.4311 | -101.4808
5.4349 | -101.4815
5.4386 | -101.4826
5.4424 | -101.4805
5.4461 | -101.4741
5.4499 | -101.4754
5.4536 | -101.4907
5.4574 | -101.5031
5.4611 | -101.5177
5.4649 | -101.542
5.4686 | -101.5724
5.4724 | -101.5985
5.4761 | -101.6138
5.4799 | -101.6171
5.4836 | -101.61
5.4874 | -101.5969
5.4911 | -101.577
5.4949 | -101.5604
5.4986 | -101.5548
5.5024 | -101.5564
5.5061 | -101.5644
5.5099 | -101.5769
5.5136 | -101.5939
5.5174 | -101.6032
5.5211 | -101.6036
5.5249 | -101.5955
5.5286 | -101.5765
5.5324 | -101.5426
5.5361 | -101.4874
5.5399 | -101.4193
5.5436 | -101.3497
5.5474 | -101.2789
5.5511 | -101.21
5.5549 | -101.1608
5.5586 | -101.1327
5.5624 | -101.1231
5.5661 | -101.1334
5.5699 | -101.1587
5.5736 | -101.1965
5.5774 | -101.2488
5.5811 | -101.3042
5.5849 | -101.3555
5.5886 | -101.4091
5.5924 | -101.4588
5.5961 | -101.5062
5.5999 | -101.549
5.6036 | -101.5951
5.6074 | -101.6394
5.6111 | -101.6814
5.6149 | -101.7266
5.6186 | -101.7823
5.6224 | -101.8416
5.6261 | -101.8923
5.6299 | -101.9405
5.6336 | -101.9802
5.6374 | -102.0079
5.6411 | -102.0284
5.6449 | -102.0487
5.6486 | -102.0639
5.6524 | -102.0675
5.6561 | -102.0687
5.6599 | -102.0629
5.6636 | -102.0474
5.6674 | -102.0342
5.6712 | -102.0229
5.6749 | -102.0085
5.6787 | -101.9863
5.6824 | -101.955
5.6862 | -101.9246
5.6899 | -101.8983
5.6937 | -101.8617
5.6974 | -101.8244
5.7012 | -101.7901
5.7049 | -101.757
5.7087 | -101.7297
5.7124 | -101.7107
5.7162 | -101.7003
5.7199 | -101.7001
5.7237 | -101.7143
5.7274 | -101.7385
5.7312 | -101.7732
5.7349 | -101.8178
5.7387 | -101.874
5.7424 | -101.9347
5.7462 | -102.004
5.7499 | -102.0687
5.7537 | -102.1229
5.7574 | -102.1641
5.7612 | -102.1921
5.7649 | -102.2208
5.7687 | -102.2358
5.7724 | -102.2405
5.7762 | -102.2382
5.7799 | -102.2317
5.7837 | -102.2245
5.7874 | -102.2056
5.7912 | -102.1836
5.7949 | -102.1664
5.7987 | -102.1478
5.8024 | -102.1236
5.8062 | -102.106
5.8099 | -102.0946
5.8137 | -102.0869
5.8174 | -102.0907
5.8212 | -102.1045
5.8249 | -102.1226
5.8287 | -102.1416
5.8324 | -102.1655
5.8362 | -102.1896
5.8399 | -102.204
5.8437 | -102.2147
5.8474 | -102.2174
5.8512 | -102.2144
5.8549 | -102.2157
5.8587 | -102.2175
5.8624 | -102.2268
5.8662 | -102.2389
5.8699 | -102.2502
5.8737 | -102.2616
5.8774 | -102.2682
5.8812 | -102.2605
5.8849 | -102.25
5.8887 | -102.2286
5.8924 | -102.1911
5.8962 | -102.1519
5.8999 | -102.1205
5.9037 | -102.0912
5.9074 | -102.0743
5.9112 | -102.0658
5.915 | -102.08
5.9187 | -102.1016
5.9225 | -102.1422
5.9262 | -102.2009
5.93 | -102.2647
5.9337 | -102.3323
5.9375 | -102.3975
5.9412 | -102.4548
5.945 | -102.5016
5.9487 | -102.514
5.9525 | -102.5402
5.9562 | -102.5527
5.96 | -102.5477
5.9637 | -102.5333
5.9675 | -102.503
5.9712 | -102.4588
5.975 | -102.4034
5.9787 | -102.3471
5.9825 | -102.2821
5.9862 | -102.209
5.99 | -102.1378
5.9937 | -102.073
5.9975 | -102.0127
6.0012 | -101.9646
6.005 | -101.9241
6.0087 | -101.9035
6.0125 | -101.8944
6.0162 | -101.8953
6.02 | -101.9159
6.0237 | -101.9466
6.0275 | -101.9811
6.0312 | -102.0254
6.035 | -102.0702
6.0387 | -102.113
6.0425 | -102.1507
6.0462 | -102.1853
6.05 | -102.2157
6.0537 | -102.2381
6.0575 | -102.2494
6.0612 | -102.2538
6.065 | -102.2522
6.0687 | -102.2375
6.0725 | -102.2183
6.0762 | -102.1971
6.08 | -102.1779
6.0837 | -102.1623
6.0875 | -102.1463
6.0912 | -102.1375
6.095 | -102.1358
6.0987 | -102.138
6.1025 | -102.141
6.1062 | -102.1523
6.11 | -102.1618
6.1137 | -102.1617
6.1175 | -102.1521
6.1212 | -102.1328
6.125 | -102.107
6.1287 | -102.0767
6.1325 | -102.049
6.1362 | -102.0205
6.14 | -101.9928
6.1437 | -101.9822
6.1475 | -101.9799
6.1512 | -101.9887
6.155 | -102.0117
6.1588 | -102.0403
6.1625 | -102.0767
6.1663 | -102.1093
6.17 | -102.1449
6.1738 | -102.1764
6.1775 | -102.1988
6.1813 | -102.2231
6.185 | -102.2449
6.1888 | -102.266
6.1925 | -102.2752
6.1963 | -102.2823
6.2 | -102.2792
6.2038 | -102.2773
6.2075 | -102.2793
6.2113 | -102.286
6.215 | -102.2959
6.2188 | -102.3088
6.2225 | -102.3293
6.2263 | -102.3558
6.23 | -102.3927
6.2338 | -102.4385
6.2375 | -102.4887
6.2413 | -102.5382
6.245 | -102.5848
6.2488 | -102.6318
6.2525 | -102.6732
6.2563 | -102.7014
6.26 | -102.7173
6.2638 | -102.7229
6.2675 | -102.7206
6.2713 | -102.7121
6.275 | -102.6984
6.2788 | -102.6801
6.2825 | -102.6548
6.2863 | -102.634
6.29 | -102.6111
6.2938 | -102.5906
6.2975 | -102.5758
6.3013 | -102.5615
6.305 | -102.5571
6.3088 | -102.5594
6.3125 | -102.5642
6.3163 | -102.5835
6.32 | -102.6141
6.3238 | -102.6483
6.3275 | -102.6871
6.3313 | -102.7327
6.335 | -102.7859
6.3388 | -102.838
6.3425 | -102.888
6.3463 | -102.943
6.35 | -102.9974
6.3538 | -103.0506
6.3575 | -103.0879
6.3613 | -103.11
6.365 | -103.1275
6.3688 | -103.1321
6.3725 | -103.1285
6.3763 | -103.1149
6.38 | -103.0882
6.3838 | -103.0563
6.3875 | -103.0148
6.3913 | -102.9634
6.395 | -102.9112
6.3988 | -102.8565
6.4026 | -102.8019
6.4063 | -102.7463
6.4101 | -102.6852
6.4138 | -102.6299
6.4176 | -102.5843
6.4213 | -102.54
6.4251 | -102.4962
6.4288 | -102.4612
6.4326 | -102.4378
6.4363 | -102.4204
6.4401 | -102.4077
6.4438 | -102.4036
6.4476 | -102.4061
6.4513 | -102.416
6.4551 | -102.4282
6.4588 | -102.4397
6.4626 | -102.4554
6.4663 | -102.4711
6.4701 | -102.4883
6.4738 | -102.502
6.4776 | -102.5011
6.4813 | -102.492
6.4851 | -102.4757
6.4888 | -102.4516
6.4926 | -102.4191
6.4963 | -102.3951
6.5001 | -102.3712
6.5038 | -102.3546
6.5076 | -102.3427
6.5113 | -102.3293
6.5151 | -102.3242
6.5188 | -102.3257
6.5226 | -102.3243
6.5263 | -102.3263
6.5301 | -102.3251
6.5338 | -102.315
6.5376 | -102.305
6.5413 | -102.2927
6.5451 | -102.2799
6.5488 | -102.2694
6.5526 | -102.2586
6.5563 | -102.2467
6.5601 | -102.2341
6.5638 | -102.222
6.5676 | -102.2118
6.5713 | -102.2086
6.5751 | -102.2112
6.5788 | -102.2229
6.5826 | -102.2396
6.5863 | -102.2643
6.5901 | -102.2996
6.5938 | -102.3476
6.5976 | -102.4016
6.6013 | -102.4675
6.6051 | -102.5348
6.6088 | -102.607
6.6126 | -102.6875
6.6163 | -102.7622
6.6201 | -102.8323
6.6238 | -102.8944
6.6276 | -102.9405
6.6313 | -102.9727
6.6351 | -102.9953
6.6388 | -103.0054
6.6426 | -103.0069
6.6463 | -103.0012
6.6501 | -102.989
6.6539 | -102.9819
6.6576 | -102.979
6.6614 | -102.9822
6.6651 | -102.9944
6.6689 | -103.018
6.6726 | -103.0464
6.6764 | -103.0857
6.6801 | -103.1307
6.6839 | -103.1807
6.6876 | -103.2261
6.6914 | -103.2641
6.6951 | -103.2938
6.6989 | -103.311
6.7026 | -103.3167
6.7064 | -103.3069
6.7101 | -103.2836
6.7139 | -103.2409
6.7176 | -103.1922
6.7214 | -103.1451
6.7251 | -103.0985
6.7289 | -103.049
6.7326 | -103.0026
6.7364 | -102.9603
6.7401 | -102.9172
6.7439 | -102.8755
6.7476 | -102.8418
6.7514 | -102.8196
6.7551 | -102.8013
6.7589 | -102.7864
6.7626 | -102.7739
6.7664 | -102.7654
6.7701 | -102.758
6.7739 | -102.7433
6.7776 | -102.7278
6.7814 | -102.7071
6.7851 | -102.6785
6.7889 | -102.6449
6.7926 | -102.6106
6.7964 | -102.5682
6.8001 | -102.533
6.8039 | -102.5003
6.8076 | -102.477
6.8114 | -102.4584
6.8151 | -102.4547
6.8189 | -102.4582
6.8226 | -102.4672
6.8264 | -102.4897
6.8301 | -102.5291
6.8339 | -102.579
6.8376 | -102.6292
6.8414 | -102.6874
6.8451 | -102.7487
6.8489 | -102.8056
6.8526 | -102.8584
6.8564 | -102.9069
6.8601 | -102.9482
6.8639 | -102.986
6.8676 | -103.0141
6.8714 | -103.0315
6.8751 | -103.0501
6.8789 | -103.0606
6.8826 | -103.0675
6.8864 | -103.0746
6.8902 | -103.0765
6.8939 | -103.0742
6.8977 | -103.075
6.9014 | -103.0657
6.9052 | -103.0493
6.9089 | -103.0391
6.9127 | -103.031
6.9164 | -103.0252
6.9202 | -103.024
6.9239 | -103.0285
6.9277 | -103.0336
6.9314 | -103.0355
6.9352 | -103.0419
6.9389 | -103.0556
6.9427 | -103.0732
6.9464 | -103.0896
6.9502 | -103.1065
6.9539 | -103.125
6.9577 | -103.1465
6.9614 | -103.1732
6.9652 | -103.2017
6.9689 | -103.2284
6.9727 | -103.2526
6.9764 | -103.2681
6.9802 | -103.2772
6.9839 | -103.2846
6.9877 | -103.2904
6.9914 | -103.2974
6.9952 | -103.3005
6.9989 | -103.3013
7.0027 | -103.3055
7.0064 | -103.3179
7.0102 | -103.3237
7.0139 | -103.3217
7.0177 | -103.3151
7.0214 | -103.3122
7.0252 | -103.3034
7.0289 | -103.2983
7.0327 | -103.3039
7.0364 | -103.3107
7.0402 | -103.3174
7.0439 | -103.3269
7.0477 | -103.3337
7.0514 | -103.3409
7.0552 | -103.3469
7.0589 | -103.351
7.0627 | -103.3542
7.0664 | -103.3556
7.0702 | -103.354
7.0739 | -103.3593
7.0777 | -103.372
7.0814 | -103.3854
7.0852 | -103.3898
7.0889 | -103.3951
7.0927 | -103.4029
7.0964 | -103.4087
7.1002 | -103.4148
7.1039 | -103.4174
7.1077 | -103.4108
7.1114 | -103.399
7.1152 | -103.3822
7.1189 | -103.3683
7.1227 | -103.3523
7.1264 | -103.3305
7.1302 | -103.3086
7.1339 | -103.2819
7.1377 | -103.2536
7.1415 | -103.2266
7.1452 | -103.2017
7.149 | -103.1767
7.1527 | -103.1561
7.1565 | -103.1351
7.1602 | -103.1173
7.164 | -103.1104
7.1677 | -103.1048
7.1715 | -103.1046
7.1752 | -103.1086
7.179 | -103.1242
7.1827 | -103.144
7.1865 | -103.164
7.1902 | -103.1912
7.194 | -103.2252
7.1977 | -103.2565
7.2015 | -103.2853
7.2052 | -103.3161
7.209 | -103.3393
7.2127 | -103.3602
7.2165 | -103.3888
7.2202 | -103.4144
7.224 | -103.4423
7.2277 | -103.4687
7.2315 | -103.4939
7.2352 | -103.5177
7.239 | -103.5428
7.2427 | -103.5659
7.2465 | -103.5874
7.2502 | -103.6041
7.254 | -103.6126
7.2577 | -103.6229
7.2615 | -103.6284
7.2652 | -103.6354
7.269 | -103.6361
7.2727 | -103.6284
7.2765 | -103.6198
7.2802 | -103.6125
7.284 | -103.6147
7.2877 | -103.6149
7.2915 | -103.6152
7.2952 | -103.6224
7.299 | -103.6302
7.3027 | -103.6384
7.3065 | -103.6357
7.3102 | -103.6351
7.314 | -103.6333
7.3177 | -103.6265
7.3215 | -103.61
7.3252 | -103.5791
7.329 | -103.5459
7.3327 | -103.5074
7.3365 | -103.4594
7.3402 | -103.4102
7.344 | -103.3625
7.3477 | -103.3121
7.3515 | -103.2668
7.3552 | -103.224
7.359 | -103.1879
7.3627 | -103.1641
7.3665 | -103.1429
7.3702 | -103.1324
7.374 | -103.1305
7.3777 | -103.1387
7.3815 | -103.1553
7.3853 | -103.1838
7.389 | -103.2143
7.3928 | -103.247
7.3965 | -103.2832
7.4003 | -103.3171
7.404 | -103.3468
7.4078 | -103.371
7.4115 | -103.3918
7.4153 | -103.4146
7.419 | -103.4229
7.4228 | -103.425
7.4265 | -103.4245
7.4303 | -103.4215
7.434 | -103.4161
7.4378 | -103.4129
7.4415 | -103.4117
7.4453 | -103.4103
7.449 | -103.4042
7.4528 | -103.4012
7.4565 | -103.4017
7.4603 | -103.3995
7.464 | -103.4024
7.4678 | -103.4109
7.4715 | -103.4256
7.4753 | -103.4475
7.479 | -103.4702
7.4828 | -103.5007
7.4865 | -103.5382
7.4903 | -103.5803
7.494 | -103.6246
7.4978 | -103.6722
7.5015 | -103.7194
7.5053 | -103.7647
7.509 | -103.8105
7.5128 | -103.8436
7.5165 | -103.8672
7.5203 | -103.8793
7.524 | -103.8804
7.5278 | -103.8769
7.5315 | -103.8613
7.5353 | -103.8393
7.539 | -103.8128
7.5428 | -103.7771
7.5465 | -103.7304
7.5503 | -103.6867
7.554 | -103.6488
7.5578 | -103.6129
7.5615 | -103.5785
7.5653 | -103.5478
7.569 | -103.5186
7.5728 | -103.4972
7.5765 | -103.4736
7.5803 | -103.4472
7.584 | -103.4213
7.5878 | -103.3993
7.5915 | -103.3769
7.5953 | -103.35
7.599 | -103.3293
7.6028 | -103.3124
7.6065 | -103.2992
7.6103 | -103.2915
7.614 | -103.291
7.6178 | -103.2946
7.6215 | -103.3069
7.6253 | -103.3269
7.6291 | -103.3511
7.6328 | -103.3794
7.6366 | -103.4122
7.6403 | -103.4431
7.6441 | -103.4741
7.6478 | -103.4976
7.6516 | -103.5172
7.6553 | -103.532
7.6591 | -103.5415
7.6628 | -103.5445
7.6666 | -103.5478
7.6703 | -103.5515
7.6741 | -103.5561
7.6778 | -103.5684
7.6816 | -103.591
7.6853 | -103.6117
7.6891 | -103.6362
7.6928 | -103.6661
7.6966 | -103.702
7.7003 | -103.7407
7.7041 | -103.7805
7.7078 | -103.8219
7.7116 | -103.8602
7.7153 | -103.8877
7.7191 | -103.9038
7.7228 | -103.91
7.7266 | -103.9094
7.7303 | -103.9027
7.7341 | -103.8945
7.7378 | -103.8723
7.7416 | -103.8524
7.7453 | -103.8358
7.7491 | -103.8184
7.7528 | -103.799
7.7566 | -103.78
7.7603 | -103.762
7.7641 | -103.7478
7.7678 | -103.7306
7.7716 | -103.7194
7.7753 | -103.7048
7.7791 | -103.6915
7.7828 | -103.6733
7.7866 | -103.6578
7.7903 | -103.6365
7.7941 | -103.6174
7.7978 | -103.6016
7.8016 | -103.5897
7.8053 | -103.5838
7.8091 | -103.5783
7.8128 | -103.5765
7.8166 | -103.5732
7.8203 | -103.5756
7.8241 | -103.5832
7.8278 | -103.5983
7.8316 | -103.6158
7.8353 | -103.6395
7.8391 | -103.6613
7.8428 | -103.6824
7.8466 | -103.7016
7.8503 | -103.7163
7.8541 | -103.7324
7.8578 | -103.7435
7.8616 | -103.7482
7.8653 | -103.7526
7.8691 | -103.7579
7.8728 | -103.7656
7.8766 | -103.7728
7.8804 | -103.7813
7.8841 | -103.7909
7.8879 | -103.8013
7.8916 | -103.8169
7.8954 | -103.8359
7.8991 | -103.8618
7.9029 | -103.8882
7.9066 | -103.9179
7.9104 | -103.9434
7.9141 | -103.9664
7.9179 | -103.9909
7.9216 | -104.0053
7.9254 | -104.0168
7.9291 | -104.016
7.9329 | -104.006
7.9366 | -103.9847
7.9404 | -103.9554
7.9441 | -103.9215
7.9479 | -103.8873
7.9516 | -103.8538
7.9554 | -103.8266
7.9591 | -103.797
7.9629 | -103.764
7.9666 | -103.7323
7.9704 | -103.703
7.9741 | -103.6782
7.9779 | -103.6584
7.9816 | -103.6403
7.9854 | -103.6283
7.9891 | -103.6187
7.9929 | -103.6113
7.9966 | -103.6066
8.0004 | -103.6053
8.0041 | -103.6061
8.0079 | -103.6125
8.0116 | -103.6207
8.0154 | -103.6311
8.0191 | -103.6463
8.0229 | -103.6679
8.0266 | -103.6914
8.0304 | -103.7139
8.0341 | -103.7332
8.0379 | -103.754
8.0416 | -103.7755
8.0454 | -103.7962
8.0491 | -103.82
8.0529 | -103.843
8.0566 | -103.8677
8.0604 | -103.8929
8.0641 | -103.9174
8.0679 | -103.9426
8.0716 | -103.9645
8.0754 | -103.9882
8.0791 | -104.016
8.0829 | -104.0434
8.0866 | -104.064
8.0904 | -104.0717
8.0941 | -104.0711
8.0979 | -104.0697
8.1016 | -104.0594
8.1054 | -104.0478
8.1091 | -104.031
8.1129 | -104.0045
8.1166 | -103.9748
8.1204 | -103.9481
8.1242 | -103.9181
8.1279 | -103.888
8.1317 | -103.8611
8.1354 | -103.8362
8.1392 | -103.8072
8.1429 | -103.783
8.1467 | -103.7637
8.1504 | -103.7503
8.1542 | -103.7442
8.1579 | -103.7398
8.1617 | -103.74
8.1654 | -103.7445
8.1692 | -103.7503
8.1729 | -103.7587
8.1767 | -103.7646
8.1804 | -103.7712
8.1842 | -103.7849
8.1879 | -103.802
8.1917 | -103.8223
8.1954 | -103.8434
8.1992 | -103.8663
8.2029 | -103.8919
8.2067 | -103.921
8.2104 | -103.9596
8.2142 | -104.0034
8.2179 | -104.0472
8.2217 | -104.0912
8.2254 | -104.1307
8.2292 | -104.1718
8.2329 | -104.211
8.2367 | -104.2352
8.2404 | -104.2557
8.2442 | -104.2755
8.2479 | -104.2882
8.2517 | -104.2958
8.2554 | -104.2953
8.2592 | -104.2913
8.2629 | -104.2868
8.2667 | -104.2722
8.2704 | -104.2553
8.2742 | -104.2322
8.2779 | -104.2059
8.2817 | -104.1798
8.2854 | -104.1535
8.2892 | -104.1287
8.2929 | -104.103
8.2967 | -104.0795
8.3004 | -104.0601
8.3042 | -104.0439
8.3079 | -104.0294
8.3117 | -104.0171
8.3154 | -104.0132
8.3192 | -104.0174
8.3229 | -104.0259
8.3267 | -104.0389
8.3304 | -104.0592
8.3342 | -104.0789
8.3379 | -104.1024
8.3417 | -104.1264
8.3454 | -104.1509
8.3492 | -104.176
8.3529 | -104.1989
8.3567 | -104.2175
8.3604 | -104.2321
8.3642 | -104.2352
8.368 | -104.2361
8.3717 | -104.2333
8.3755 | -104.2209
8.3792 | -104.2097
8.383 | -104.1973
8.3867 | -104.1813
8.3905 | -104.1695
8.3942 | -104.1552
8.398 | -104.1445
8.4017 | -104.132
8.4055 | -104.1176
8.4092 | -104.101
8.413 | -104.0922
8.4167 | -104.0832
8.4205 | -104.0703
8.4242 | -104.0646
8.428 | -104.06
8.4317 | -104.054
8.4355 | -104.0489
8.4392 | -104.0451
8.443 | -104.0403
8.4467 | -104.0409
8.4505 | -104.0452
8.4542 | -104.0469
8.458 | -104.051
8.4617 | -104.057
8.4655 | -104.0624
8.4692 | -104.0733
8.473 | -104.0906
8.4767 | -104.1068
8.4805 | -104.1271
8.4842 | -104.1519
8.488 | -104.185
8.4917 | -104.2204
8.4955 | -104.2582
8.4992 | -104.2982
8.503 | -104.3355
8.5067 | -104.3747
8.5105 | -104.4147
8.5142 | -104.448
8.518 | -104.4787
8.5217 | -104.4986
8.5255 | -104.5155
8.5292 | -104.5294
8.533 | -104.5395
8.5367 | -104.5392
8.5405 | -104.5348
8.5442 | -104.5258
8.548 | -104.5068
8.5517 | -104.4767
8.5555 | -104.4431
8.5592 | -104.4103
8.563 | -104.3648
8.5667 | -104.3123
8.5705 | -104.2505
8.5742 | -104.1839
8.578 | -104.1185
8.5817 | -104.0536
8.5855 | -103.9909
8.5892 | -103.938
8.593 | -103.8889
8.5967 | -103.8467
8.6005 | -103.8124
8.6042 | -103.7919
8.608 | -103.7821
8.6117 | -103.782
8.6155 | -103.7892
8.6193 | -103.8101
8.623 | -103.8456
8.6268 | -103.889
8.6305 | -103.9366
8.6343 | -103.9914
8.638 | -104.048
8.6418 | -104.109
8.6455 | -104.1729
8.6493 | -104.2367
8.653 | -104.2956
8.6568 | -104.3438
8.6605 | -104.3913
8.6643 | -104.434
8.668 | -104.4763
8.6718 | -104.5175
8.6755 | -104.5536
8.6793 | -104.5894
8.683 | -104.6177
8.6868 | -104.6358
8.6905 | -104.6507
8.6943 | -104.6578
8.698 | -104.6668
8.7018 | -104.6694
8.7055 | -104.6664
8.7093 | -104.6633
8.713 | -104.6513
8.7168 | -104.6324
8.7205 | -104.6097
8.7243 | -104.5814
8.728 | -104.5509
8.7318 | -104.5112
8.7355 | -104.4658
8.7393 | -104.4126
8.743 | -104.3603
8.7468 | -104.3087
8.7505 | -104.2542
8.7543 | -104.2065
8.758 | -104.1684
8.7618 | -104.1362
8.7655 | -104.112
8.7693 | -104.0962
8.773 | -104.0924
8.7768 | -104.0922
8.7805 | -104.1011
8.7843 | -104.1166
8.788 | -104.1384
8.7918 | -104.1707
8.7955 | -104.211
8.7993 | -104.2506
8.803 | -104.2884
8.8068 | -104.3284
8.8105 | -104.3715
8.8143 | -104.4137
8.818 | -104.4541
8.8218 | -104.4965
8.8255 | -104.5393
8.8293 | -104.5752
8.833 | -104.6097
8.8368 | -104.644
8.8405 | -104.6736
8.8443 | -104.7024
8.848 | -104.7246
8.8518 | -104.7413
8.8555 | -104.7547
8.8593 | -104.7631
8.863 | -104.7681
8.8668 | -104.7684
8.8706 | -104.7665
8.8743 | -104.7577
8.8781 | -104.7441
8.8818 | -104.734
8.8856 | -104.7161
8.8893 | -104.6954
8.8931 | -104.6808
8.8968 | -104.67
8.9006 | -104.6566
8.9043 | -104.6467
8.9081 | -104.6415
8.9118 | -104.6367
8.9156 | -104.6347
8.9193 | -104.6358
8.9231 | -104.6369
8.9268 | -104.6379
8.9306 | -104.6424
8.9343 | -104.6514
8.9381 | -104.6603
8.9418 | -104.6717
8.9456 | -104.6844
8.9493 | -104.6982
8.9531 | -104.7154
8.9568 | -104.7292
8.9606 | -104.7405
8.9643 | -104.7543
8.9681 | -104.7635
8.9718 | -104.7723
8.9756 | -104.7849
8.9793 | -104.8045
8.9831 | -104.8272
8.9868 | -104.8473
8.9906 | -104.8675
8.9943 | -104.8898
8.9981 | -104.9136
9.0018 | -104.9333
9.0056 | -104.9541
9.0093 | -104.9728
9.0131 | -104.9878
9.0168 | -105.0072
9.0206 | -105.0245
9.0243 | -105.037
9.0281 | -105.0472
9.0318 | -105.0587
9.0356 | -105.0635
9.0393 | -105.0638
9.0431 | -105.0613
9.0468 | -105.0562
9.0506 | -105.0489
9.0543 | -105.0355
9.0581 | -105.0164
9.0618 | -104.9934
9.0656 | -104.9678
9.0693 | -104.9403
9.0731 | -104.9036
9.0768 | -104.8674
9.0806 | -104.8338
9.0843 | -104.7985
9.0881 | -104.7633
9.0918 | -104.7313
9.0956 | -104.6972
9.0993 | -104.6677
9.1031 | -104.6433
9.1068 | -104.6203
9.1106 | -104.5966
9.1144 | -104.5718
9.1181 | -104.5551
9.1219 | -104.5432
9.1256 | -104.533
9.1294 | -104.5286
9.1331 | -104.5314
9.1369 | -104.5347
9.1406 | -104.5415
9.1444 | -104.5492
9.1481 | -104.5577
9.1519 | -104.5652
9.1556 | -104.5733
9.1594 | -104.5854
9.1631 | -104.5968
9.1669 | -104.6139
9.1706 | -104.6366
9.1744 | -104.6591
9.1781 | -104.6777
9.1819 | -104.6995
9.1856 | -104.7266
9.1894 | -104.7528
9.1931 | -104.7811
9.1969 | -104.8086
9.2006 | -104.8331
9.2044 | -104.8565
9.2081 | -104.8849
9.2119 | -104.9119
9.2156 | -104.9326
9.2194 | -104.9481
9.2231 | -104.9676
9.2269 | -104.9846
9.2306 | -104.9987
9.2344 | -105.0078
9.2381 | -105.0128
9.2419 | -105.0192
9.2456 | -105.0267
9.2494 | -105.0356
9.2531 | -105.0401
9.2569 | -105.0376
9.2606 | -105.0294
9.2644 | -105.0183
9.2681 | -105.0069
9.2719 | -104.9998
9.2756 | -104.9909
9.2794 | -104.975
9.2831 | -104.9613
9.2869 | -104.9478
9.2906 | -104.929
9.2944 | -104.9104
9.2981 | -104.895
9.3019 | -104.8786
9.3056 | -104.8611
9.3094 | -104.8427
9.3131 | -104.8235
9.3169 | -104.8047
9.3206 | -104.7898
9.3244 | -104.7732
9.3281 | -104.7583
9.3319 | -104.7404
9.3356 | -104.7262
9.3394 | -104.7144
9.3431 | -104.7031
9.3469 | -104.6962
9.3506 | -104.6913
9.3544 | -104.6876
9.3582 | -104.6858
9.3619 | -104.6822
9.3657 | -104.6796
9.3694 | -104.6788
9.3732 | -104.6823
9.3769 | -104.6843
9.3807 | -104.6891
9.3844 | -104.6958
9.3882 | -104.705
9.3919 | -104.7134
9.3957 | -104.7292
9.3994 | -104.7489
9.4032 | -104.769
9.4069 | -104.786
9.4107 | -104.8
9.4144 | -104.819
9.4182 | -104.8457
9.4219 | -104.8678
9.4257 | -104.8868
9.4294 | -104.9075
9.4332 | -104.9226
9.4369 | -104.9294
9.4407 | -104.9347
9.4444 | -104.9405
9.4482 | -104.9453
9.4519 | -104.9501
9.4557 | -104.9486
9.4594 | -104.9447
9.4632 | -104.942
9.4669 | -104.9339
9.4707 | -104.9298
9.4744 | -104.9232
9.4782 | -104.9112
9.4819 | -104.8959
9.4857 | -104.885
9.4894 | -104.8748
9.4932 | -104.8628
9.4969 | -104.8513
9.5007 | -104.8406
9.5044 | -104.8282
9.5082 | -104.8151
9.5119 | -104.8009
9.5157 | -104.7888
9.5194 | -104.7696
9.5232 | -104.7523
9.5269 | -104.7348
9.5307 | -104.7225
9.5344 | -104.7106
9.5382 | -104.6995
9.5419 | -104.6889
9.5457 | -104.6819
9.5494 | -104.6769
9.5532 | -104.6726
9.5569 | -104.6682
9.5607 | -104.6679
9.5644 | -104.6738
9.5682 | -104.6812
9.5719 | -104.692
9.5757 | -104.7075
9.5794 | -104.7244
9.5832 | -104.7429
9.5869 | -104.7641
9.5907 | -104.784
9.5944 | -104.8083
9.5982 | -104.8319
9.602 | -104.8586
9.6057 | -104.882
9.6095 | -104.9007
9.6132 | -104.9168
9.617 | -104.931
9.6207 | -104.9437
9.6245 | -104.9481
9.6282 | -104.9529
9.632 | -104.9612
9.6357 | -104.9661
9.6395 | -104.9671
9.6432 | -104.9692
9.647 | -104.9698
9.6507 | -104.9687
9.6545 | -104.9704
9.6582 | -104.9766
9.662 | -104.9809
9.6657 | -104.9881
9.6695 | -104.9959
9.6732 | -105.0009
9.677 | -105.0036
9.6807 | -105.0086
9.6845 | -105.0127
9.6882 | -105.0147
9.692 | -105.0151
9.6957 | -105.0215
9.6995 | -105.0252
9.7032 | -105.026
9.707 | -105.0265
9.7107 | -105.0299
9.7145 | -105.0372
9.7182 | -105.0451
9.722 | -105.0509
9.7257 | -105.0553
9.7295 | -105.0625
9.7332 | -105.0662
9.737 | -105.058
9.7407 | -105.0544
9.7445 | -105.0493
9.7482 | -105.0458
9.752 | -105.0431
9.7557 | -105.0373
9.7595 | -105.0311
9.7632 | -105.0205
9.767 | -105.0083
9.7707 | -104.994
9.7745 | -104.9792
9.7782 | -104.9687
9.782 | -104.9628
9.7857 | -104.9569
9.7895 | -104.9535
9.7932 | -104.9513
9.797 | -104.9459
9.8007 | -104.9422
9.8045 | -104.9433
9.8082 | -104.9449
9.812 | -104.9454
9.8157 | -104.9524
9.8195 | -104.9619
9.8232 | -104.9708
9.827 | -104.9786
9.8307 | -104.9841
9.8345 | -104.9892
9.8382 | -104.9893
9.842 | -104.9821
9.8458 | -104.9705
9.8495 | -104.9578
9.8533 | -104.9424
9.857 | -104.9188
9.8608 | -104.8911
9.8645 | -104.8671
9.8683 | -104.8457
9.872 | -104.8207
9.8758 | -104.7974
9.8795 | -104.7803
9.8833 | -104.7608
9.887 | -104.7419
9.8908 | -104.7279
9.8945 | -104.719
9.8983 | -104.71
9.902 | -104.7055
9.9058 | -104.7059
9.9095 | -104.7101
9.9133 | -104.7145
9.917 | -104.7221
9.9208 | -104.7311
9.9245 | -104.7382
9.9283 | -104.7432
9.932 | -104.7491
9.9358 | -104.757
9.9395 | -104.7686
9.9433 | -104.7851
9.947 | -104.8022
9.9508 | -104.8156
9.9545 | -104.8249
9.9583 | -104.8362
9.962 | -104.8487
9.9658 | -104.8656
9.9695 | -104.8805
9.9733 | -104.8933
9.977 | -104.9082
9.9808 | -104.9187
9.9845 | -104.923
9.9883 | -104.9316
9.992 | -104.9361
9.9958 | -104.9338
9.9995 | -104.9247
10.0033 | -104.9187
10.007 | -104.9117
10.0108 | -104.8986
10.0145 | -104.8813
10.0183 | -104.861
10.022 | -104.8444
10.0258 | -104.8273
10.0295 | -104.8155
10.0333 | -104.8031
10.037 | -104.7886
10.0408 | -104.7781
10.0445 | -104.7736
10.0483 | -104.7723
10.052 | -104.7734
10.0558 | -104.7799
10.0595 | -104.7897
10.0633 | -104.8025
10.067 | -104.8184
10.0708 | -104.8395
10.0745 | -104.8629
10.0783 | -104.8866
10.082 | -104.9103
10.0858 | -104.9313
10.0895 | -104.9509
10.0933 | -104.971
10.0971 | -104.9894
10.1008 | -105.0087
10.1046 | -105.0247
10.1083 | -105.0324
10.1121 | -105.0408
10.1158 | -105.0505
10.1196 | -105.0587
10.1233 | -105.0635
10.1271 | -105.0677
10.1308 | -105.0759
10.1346 | -105.0779
10.1383 | -105.0796
10.1421 | -105.0782
10.1458 | -105.0759
10.1496 | -105.0689
10.1533 | -105.0618
10.1571 | -105.0547
10.1608 | -105.0482
10.1646 | -105.0421
10.1683 | -105.0319
10.1721 | -105.0211
10.1758 | -105.0104
10.1796 | -105.0037
10.1833 | -105.001
10.1871 | -104.995
10.1908 | -104.9885
10.1946 | -104.9859
10.1983 | -104.9842
10.2021 | -104.9818
10.2058 | -104.9814
10.2096 | -104.982
10.2133 | -104.9809
10.2171 | -104.9784
10.2208 | -104.9759
10.2246 | -104.9761
10.2283 | -104.9744
10.2321 | -104.9687
10.2358 | -104.9639
10.2396 | -104.9598
10.2433 | -104.9577
10.2471 | -104.9516
10.2508 | -104.9425
10.2546 | -104.933
10.2583 | -104.9146
10.2621 | -104.8951
10.2658 | -104.8788
10.2696 | -104.8606
10.2733 | -104.8419
10.2771 | -104.8207
10.2808 | -104.8073
10.2846 | -104.7954
10.2883 | -104.7876
10.2921 | -104.7826
10.2958 | -104.782
10.2996 | -104.7835
10.3033 | -104.7911
10.3071 | -104.8054
10.3108 | -104.8259
10.3146 | -104.8515
10.3183 | -104.8808
10.3221 | -104.9118
10.3258 | -104.9457
10.3296 | -104.9805
10.3333 | -105.0137
10.3371 | -105.0502
10.3408 | -105.09
10.3446 | -105.1253
10.3484 | -105.1571
10.3521 | -105.1853
10.3559 | -105.2128
10.3596 | -105.2363
10.3634 | -105.2549
10.3671 | -105.2704
10.3709 | -105.285
10.3746 | -105.2935
10.3784 | -105.2976
10.3821 | -105.2995
10.3859 | -105.2971
10.3896 | -105.2904
10.3934 | -105.2797
10.3971 | -105.2708
10.4009 | -105.2618
10.4046 | -105.2541
10.4084 | -105.2467
10.4121 | -105.2378
10.4159 | -105.2296
10.4196 | -105.2253
10.4234 | -105.2239
10.4271 | -105.2235
10.4309 | -105.2229
10.4346 | -105.2254
10.4384 | -105.2247
10.4421 | -105.2245
10.4459 | -105.2278
10.4496 | -105.228
10.4534 | -105.2297
10.4571 | -105.2268
10.4609 | -105.2224
10.4646 | -105.2121
10.4684 | -105.2044
10.4721 | -105.1959
10.4759 | -105.1876
10.4796 | -105.1785
10.4834 | -105.1707
10.4871 | -105.1612
10.4909 | -105.149
10.4946 | -105.136
10.4984 | -105.1267
10.5021 | -105.1202
10.5059 | -105.1184
10.5096 | -105.115
10.5134 | -105.114
10.5171 | -105.1159
10.5209 | -105.1158
10.5246 | -105.118
10.5284 | -105.1246
10.5321 | -105.1349
10.5359 | -105.1506
10.5396 | -105.1627
10.5434 | -105.1758
10.5471 | -105.193
10.5509 | -105.2113
10.5546 | -105.2304
10.5584 | -105.2506
10.5621 | -105.272
10.5659 | -105.2919
10.5696 | -105.3106
10.5734 | -105.3264
10.5771 | -105.3425
10.5809 | -105.3616
10.5846 | -105.3816
10.5884 | -105.399
10.5922 | -105.4145
10.5959 | -105.4237
10.5997 | -105.4289
10.6034 | -105.4387
10.6072 | -105.4483
10.6109 | -105.4571
10.6147 | -105.4627
10.6184 | -105.4694
10.6222 | -105.4687
10.6259 | -105.4612
10.6297 | -105.4569
10.6334 | -105.4502
10.6372 | -105.4365
10.6409 | -105.4256
10.6447 | -105.4123
10.6484 | -105.3964
10.6522 | -105.3783
10.6559 | -105.3591
10.6597 | -105.3402
10.6634 | -105.3236
10.6672 | -105.3104
10.6709 | -105.301
10.6747 | -105.2963
10.6784 | -105.2926
10.6822 | -105.2873
10.6859 | -105.2802
10.6897 | -105.2767
10.6934 | -105.2852
10.6972 | -105.2968
10.7009 | -105.3085
10.7047 | -105.3217
10.7084 | -105.337
10.7122 | -105.3532
10.7159 | -105.3711
10.7197 | -105.3902
10.7234 | -105.4094
10.7272 | -105.43
10.7309 | -105.4499
10.7347 | -105.4652
10.7384 | -105.4811
10.7422 | -105.496
10.7459 | -105.5121
10.7497 | -105.5302
10.7534 | -105.548
10.7572 | -105.5575
10.7609 | -105.5611
10.7647 | -105.5629
10.7684 | -105.5658
10.7722 | -105.5667
10.7759 | -105.5703
10.7797 | -105.5696
10.7834 | -105.5608
10.7872 | -105.5533
10.7909 | -105.548
10.7947 | -105.5327
10.7984 | -105.5128
10.8022 | -105.4881
10.8059 | -105.4602
10.8097 | -105.4322
10.8134 | -105.4001
10.8172 | -105.3664
10.8209 | -105.3292
10.8247 | -105.2951
10.8285 | -105.2621
10.8322 | -105.2301
10.836 | -105.1989
10.8397 | -105.1707
10.8435 | -105.1501
10.8472 | -105.1346
10.851 | -105.1179
10.8547 | -105.1013
10.8585 | -105.0874
10.8622 | -105.0786
10.866 | -105.0759
10.8697 | -105.0759
10.8735 | -105.0762
10.8772 | -105.0811
10.881 | -105.0902
10.8847 | -105.0963
10.8885 | -105.1035
10.8922 | -105.1143
10.896 | -105.1238
10.8997 | -105.1341
10.9035 | -105.1389
10.9072 | -105.1404
10.911 | -105.1413
10.9147 | -105.1447
10.9185 | -105.1467
10.9222 | -105.1461
10.926 | -105.1446
10.9297 | -105.1421
10.9335 | -105.1313
10.9372 | -105.1246
10.941 | -105.1197
10.9447 | -105.1101
10.9485 | -105.1008
10.9522 | -105.0941
10.956 | -105.0894
10.9597 | -105.0841
10.9635 | -105.0779
10.9672 | -105.0691
10.971 | -105.0594
10.9747 | -105.0488
10.9785 | -105.0418
10.9822 | -105.0349
10.986 | -105.0283
10.9897 | -105.0213
10.9935 | -105.0182
10.9972 | -105.0184
11.001 | -105.0197
11.0047 | -105.02
11.0085 | -105.0193
11.0122 | -105.0207
11.016 | -105.0289
11.0197 | -105.0425
11.0235 | -105.0555
11.0272 | -105.0633
11.031 | -105.0729
11.0347 | -105.0901
11.0385 | -105.1057
11.0422 | -105.1165
11.046 | -105.1302
11.0497 | -105.1418
11.0535 | -105.1516
11.0572 | -105.1619
11.061 | -105.1735
11.0647 | -105.1865
11.0685 | -105.2007
11.0722 | -105.2123
11.076 | -105.2207
11.0798 | -105.2311
11.0835 | -105.2382
11.0873 | -105.2453
11.091 | -105.2536
11.0948 | -105.2666
11.0985 | -105.2797
11.1023 | -105.2902
11.106 | -105.3024
11.1098 | -105.3144
11.1135 | -105.3318
11.1173 | -105.3491
11.121 | -105.3683
11.1248 | -105.3889
11.1285 | -105.4113
11.1323 | -105.4317
11.136 | -105.4496
11.1398 | -105.4625
11.1435 | -105.4741
11.1473 | -105.4828
11.151 | -105.4911
11.1548 | -105.4925
11.1585 | -105.4924
11.1623 | -105.4905
11.166 | -105.4845
11.1698 | -105.4752
11.1735 | -105.4645
11.1773 | -105.4535
11.181 | -105.4413
11.1848 | -105.427
11.1885 | -105.4186
11.1923 | -105.414
11.196 | -105.4028
11.1998 | -105.3921
11.2035 | -105.3901
11.2073 | -105.3909
11.211 | -105.3891
11.2148 | -105.3905
11.2185 | -105.3951
11.2223 | -105.401
11.226 | -105.4085
11.2298 | -105.4143
11.2335 | -105.4236
11.2373 | -105.4355
11.241 | -105.4489
11.2448 | -105.4656
11.2485 | -105.4785
11.2523 | -105.4929
11.256 | -105.5082
11.2598 | -105.5234
11.2635 | -105.5398
11.2673 | -105.5557
11.271 | -105.5705
11.2748 | -105.5836
11.2785 | -105.5957
11.2823 | -105.6099
11.286 | -105.6252
11.2898 | -105.636
11.2935 | -105.6436
11.2973 | -105.6532
11.301 | -105.6603
11.3048 | -105.6655
11.3085 | -105.6704
11.3123 | -105.6692
11.316 | -105.6684
11.3198 | -105.664
11.3235 | -105.6593
11.3273 | -105.6509
11.3311 | -105.6365
11.3348 | -105.6198
11.3386 | -105.5969
11.3423 | -105.5723
11.3461 | -105.5463
11.3498 | -105.5155
11.3536 | -105.484
11.3573 | -105.4554
11.3611 | -105.4314
11.3648 | -105.4097
11.3686 | -105.3846
11.3723 | -105.3659
11.3761 | -105.3556
11.3798 | -105.3429
11.3836 | -105.3312
11.3873 | -105.321
11.3911 | -105.3215
11.3948 | -105.3252
11.3986 | -105.3222
11.4023 | -105.3238
11.4061 | -105.3248
11.4098 | -105.3296
11.4136 | -105.3371
11.4173 | -105.3423
11.4211 | -105.3478
11.4248 | -105.349
11.4286 | -105.3467
11.4323 | -105.3434
11.4361 | -105.3371
11.4398 | -105.3277
11.4436 | -105.314
11.4473 | -105.2921
11.4511 | -105.2673
11.4548 | -105.2448
11.4586 | -105.2235
11.4623 | -105.2019
11.4661 | -105.1789
11.4698 | -105.1539
11.4736 | -105.125
11.4773 | -105.0991
11.4811 | -105.0748
11.4848 | -105.0527
11.4886 | -105.0337
11.4923 | -105.0224
11.4961 | -105.0147
11.4998 | -105.0098
11.5036 | -105.0121
11.5073 | -105.0186
11.5111 | -105.0314
11.5148 | -105.0487
11.5186 | -105.0661
11.5223 | -105.0883
11.5261 | -105.1139
11.5298 | -105.1383
11.5336 | -105.1655
11.5373 | -105.1976
11.5411 | -105.2277
11.5448 | -105.258
11.5486 | -105.2905
11.5523 | -105.3229
11.5561 | -105.3547
11.5598 | -105.3804
11.5636 | -105.4044
11.5673 | -105.4291
11.5711 | -105.448
11.5749 | -105.4637
11.5786 | -105.474
11.5824 | -105.4802
11.5861 | -105.4821
11.5899 | -105.4788
11.5936 | -105.4746
11.5974 | -105.4714
11.6011 | -105.4657
11.6049 | -105.4586
11.6086 | -105.4495
11.6124 | -105.4407
11.6161 | -105.4328
11.6199 | -105.4247
11.6236 | -105.4134
11.6274 | -105.4035
11.6311 | -105.3911
11.6349 | -105.3814
11.6386 | -105.3757
11.6424 | -105.3767
11.6461 | -105.3768
11.6499 | -105.377
11.6536 | -105.3761
11.6574 | -105.3756
11.6611 | -105.3764
11.6649 | -105.3755
11.6686 | -105.3738
11.6724 | -105.371
11.6761 | -105.3691
11.6799 | -105.3682
11.6836 | -105.3641
11.6874 | -105.3606
11.6911 | -105.3583
11.6949 | -105.3587
11.6986 | -105.3619
11.7024 | -105.3628
11.7061 | -105.364
11.7099 | -105.3687
11.7136 | -105.3745
11.7174 | -105.3837
11.7211 | -105.3918
11.7249 | -105.4017
11.7286 | -105.4141
11.7324 | -105.4294
11.7361 | -105.4505
11.7399 | -105.4725
11.7436 | -105.4921
11.7474 | -105.5089
11.7511 | -105.5288
11.7549 | -105.5492
11.7586 | -105.565
11.7624 | -105.5777
11.7661 | -105.5917
11.7699 | -105.6051
11.7736 | -105.6182
11.7774 | -105.6283
11.7811 | -105.6371
11.7849 | -105.6439
11.7886 | -105.6449
11.7924 | -105.6394
11.7961 | -105.63
11.7999 | -105.6189
11.8036 | -105.6053
11.8074 | -105.589
11.8111 | -105.5682
11.8149 | -105.5497
11.8187 | -105.53
11.8224 | -105.5134
11.8262 | -105.497
11.8299 | -105.4799
11.8337 | -105.4703
11.8374 | -105.464
11.8412 | -105.4548
11.8449 | -105.4453
11.8487 | -105.4405
11.8524 | -105.4451
11.8562 | -105.4513
11.8599 | -105.4562
11.8637 | -105.4613
11.8674 | -105.4682
11.8712 | -105.4778
11.8749 | -105.487
11.8787 | -105.4987
11.8824 | -105.5118
11.8862 | -105.5227
11.8899 | -105.536
11.8937 | -105.5489
11.8974 | -105.5601
11.9012 | -105.576
11.9049 | -105.5936
11.9087 | -105.609
11.9124 | -105.6221
11.9162 | -105.6375
11.9199 | -105.6564
11.9237 | -105.6745
11.9274 | -105.6954
11.9312 | -105.7178
11.9349 | -105.7363
11.9387 | -105.7461
11.9424 | -105.7521
11.9462 | -105.7591
11.9499 | -105.7677
11.9537 | -105.7746
11.9574 | -105.778
11.9612 | -105.7811
11.9649 | -105.7862
11.9687 | -105.7852
11.9724 | -105.7818
11.9762 | -105.7743
11.9799 | -105.7648
11.9837 | -105.7543
11.9874 | -105.7404
11.9912 | -105.7237
11.9949 | -105.7072
11.9987 | -105.6906
12.0024 | -105.6763
12.0062 | -105.66
12.0099 | -105.6413
12.0137 | -105.625
12.0174 | -105.6125
12.0212 | -105.6018
12.0249 | -105.5892
12.0287 | -105.5806
12.0324 | -105.5785
12.0362 | -105.5778
12.0399 | -105.5809
12.0437 | -105.5835
12.0474 | -105.587
12.0512 | -105.5895
12.0549 | -105.5945
12.0587 | -105.6068
12.0625 | -105.6213
12.0662 | -105.6355
12.07 | -105.6516
12.0737 | -105.6649
12.0775 | -105.6755
12.0812 | -105.6838
12.085 | -105.6932
12.0887 | -105.7049
12.0925 | -105.719
12.0962 | -105.7312
12.1 | -105.7418
12.1037 | -105.7491
12.1075 | -105.7585
12.1112 | -105.7658
12.115 | -105.7701
12.1187 | -105.7772
12.1225 | -105.7758
12.1262 | -105.7695
12.13 | -105.7619
12.1337 | -105.7577
12.1375 | -105.7505
12.1412 | -105.7404
12.145 | -105.7295
12.1487 | -105.7193
12.1525 | -105.7106
12.1562 | -105.6973
12.16 | -105.6827
12.1637 | -105.6701
12.1675 | -105.6573
12.1712 | -105.6423
12.175 | -105.6298
12.1787 | -105.6162
12.1825 | -105.605
12.1862 | -105.5984
12.19 | -105.5927
12.1937 | -105.5835
12.1975 | -105.5724
12.2012 | -105.5657
12.205 | -105.5626
12.2087 | -105.5594
12.2125 | -105.5583
12.2162 | -105.557
12.22 | -105.5568
12.2237 | -105.5603
12.2275 | -105.5659
12.2312 | -105.5736
12.235 | -105.5792
12.2387 | -105.5851
12.2425 | -105.5931
12.2462 | -105.599
12.25 | -105.6051
12.2537 | -105.6159
12.2575 | -105.6275
12.2612 | -105.6364
12.265 | -105.6442
12.2687 | -105.6498
12.2725 | -105.6556
12.2762 | -105.6635
12.28 | -105.6756
12.2837 | -105.684
12.2875 | -105.6957
12.2912 | -105.7106
12.295 | -105.7197
12.2987 | -105.7306
12.3025 | -105.7409
12.3063 | -105.7451
12.31 | -105.7441
12.3138 | -105.7425
12.3175 | -105.7416
12.3213 | -105.741
12.325 | -105.7338
12.3288 | -105.7256
12.3325 | -105.724
12.3363 | -105.7195
12.34 | -105.7129
12.3438 | -105.7085
12.3475 | -105.7022
12.3513 | -105.6953
12.355 | -105.6875
12.3588 | -105.6783
12.3625 | -105.6675
12.3663 | -105.6567
12.37 | -105.6474
12.3738 | -105.6375
12.3775 | -105.6283
12.3813 | -105.6191
12.385 | -105.6111
12.3888 | -105.604
12.3925 | -105.5986
12.3963 | -105.595
12.4 | -105.5929
12.4038 | -105.5913
12.4075 | -105.591
12.4113 | -105.5939
12.415 | -105.5979
12.4188 | -105.6052
12.4225 | -105.6137
12.4263 | -105.6249
12.43 | -105.6329
12.4338 | -105.6408
12.4375 | -105.6552
12.4413 | -105.6697
12.445 | -105.6782
12.4488 | -105.6872
12.4525 | -105.694
12.4563 | -105.7005
12.46 | -105.7058
12.4638 | -105.7111
12.4675 | -105.7151
12.4713 | -105.7185
12.475 | -105.7218
12.4788 | -105.7259
12.4825 | -105.7288
12.4863 | -105.7317
12.49 | -105.7349
12.4938 | -105.7347
12.4975 | -105.7312
12.5013 | -105.7269
12.505 | -105.7219
12.5088 | -105.7161
12.5125 | -105.7087
12.5163 | -105.7055
12.52 | -105.7006
12.5238 | -105.6919
12.5275 | -105.6837
12.5313 | -105.6769
12.535 | -105.6718
12.5388 | -105.6668
12.5425 | -105.6629
12.5463 | -105.6563
12.55 | -105.6513
12.5538 | -105.647
12.5576 | -105.6418
12.5613 | -105.6355
12.5651 | -105.6304
12.5688 | -105.6272
12.5726 | -105.6259
12.5763 | -105.6253
12.5801 | -105.6217
12.5838 | -105.6215
12.5876 | -105.6208
12.5913 | -105.6185
12.5951 | -105.6185
12.5988 | -105.6157
12.6026 | -105.6154
12.6063 | -105.6167
12.6101 | -105.6188
12.6138 | -105.6144
12.6176 | -105.6102
12.6213 | -105.6092
12.6251 | -105.6097
12.6288 | -105.6113
12.6326 | -105.6106
12.6363 | -105.6095
12.6401 | -105.6057
12.6438 | -105.6032
12.6476 | -105.602
12.6513 | -105.6018
12.6551 | -105.5999
12.6588 | -105.5983
12.6626 | -105.5959
12.6663 | -105.5949
12.6701 | -105.5951
12.6738 | -105.5945
12.6776 | -105.5939
12.6813 | -105.5956
12.6851 | -105.5977
12.6888 | -105.6065
12.6926 | -105.6165
12.6963 | -105.6272
12.7001 | -105.638
12.7038 | -105.6471
12.7076 | -105.6649
12.7113 | -105.6869
12.7151 | -105.7073
12.7188 | -105.7268
12.7226 | -105.745
12.7263 | -105.762
12.7301 | -105.7819
12.7338 | -105.799
12.7376 | -105.8141
12.7413 | -105.8252
12.7451 | -105.8328
12.7488 | -105.8401
12.7526 | -105.8457
12.7563 | -105.8478
12.7601 | -105.849
12.7638 | -105.8452
12.7676 | -105.8376
12.7713 | -105.8305
12.7751 | -105.8227
12.7788 | -105.8128
12.7826 | -105.8052
12.7863 | -105.7964
12.7901 | -105.7833
12.7938 | -105.7699
12.7976 | -105.7626
12.8013 | -105.7482
12.8051 | -105.7296
12.8089 | -105.7124
12.8126 | -105.7
12.8164 | -105.6859
12.8201 | -105.6715
12.8239 | -105.6579
12.8276 | -105.6443
12.8314 | -105.6315
12.8351 | -105.6179
12.8389 | -105.6068
12.8426 | -105.5974
12.8464 | -105.5878
12.8501 | -105.5787
12.8539 | -105.574
12.8576 | -105.5679
12.8614 | -105.5651
12.8651 | -105.5626
12.8689 | -105.5597
12.8726 | -105.5548
12.8764 | -105.554
12.8801 | -105.5549
12.8839 | -105.5559
12.8876 | -105.5643
12.8914 | -105.5697
12.8951 | -105.5712
12.8989 | -105.576
12.9026 | -105.5845
12.9064 | -105.5979
12.9101 | -105.6102
12.9139 | -105.6192
12.9176 | -105.6298
12.9214 | -105.6414
12.9251 | -105.6531
12.9289 | -105.6663
12.9326 | -105.6784
12.9364 | -105.6904
12.9401 | -105.7017
12.9439 | -105.7122
12.9476 | -105.7254
12.9514 | -105.7358
12.9551 | -105.7421
12.9589 | -105.7489
12.9626 | -105.7565
12.9664 | -105.7626
12.9701 | -105.7648
12.9739 | -105.7661
12.9776 | -105.7652
12.9814 | -105.7637
12.9851 | -105.7532
12.9889 | -105.7522
12.9926 | -105.7522
12.9964 | -105.7469
13.0001 | -105.7394
13.0039 | -105.7352
13.0076 | -105.7287
13.0114 | -105.7198
13.0151 | -105.7076
13.0189 | -105.697
13.0226 | -105.6868
13.0264 | -105.6751
13.0301 | -105.6641
13.0339 | -105.6567
13.0376 | -105.6485
13.0414 | -105.6408
13.0451 | -105.6327
13.0489 | -105.6264
13.0527 | -105.6186
13.0564 | -105.613
13.0602 | -105.6092
13.0639 | -105.6035
13.0677 | -105.5995
13.0714 | -105.5996
13.0752 | -105.6012
13.0789 | -105.6047
13.0827 | -105.611
13.0864 | -105.617
13.0902 | -105.6254
13.0939 | -105.6344
13.0977 | -105.6452
13.1014 | -105.6587
13.1052 | -105.6679
13.1089 | -105.6743
13.1127 | -105.687
13.1164 | -105.7009
13.1202 | -105.7124
13.1239 | -105.7247
13.1277 | -105.7386
13.1314 | -105.7515
13.1352 | -105.7611
13.1389 | -105.7723
13.1427 | -105.784
13.1464 | -105.7966
13.1502 | -105.8051
13.1539 | -105.8134
13.1577 | -105.8213
13.1614 | -105.8308
13.1652 | -105.844
13.1689 | -105.8535
13.1727 | -105.8622
13.1764 | -105.8694
13.1802 | -105.8726
13.1839 | -105.8754
13.1877 | -105.8808
13.1914 | -105.8806
13.1952 | -105.8805
13.1989 | -105.8817
13.2027 | -105.8833
13.2064 | -105.8857
13.2102 | -105.8849
13.2139 | -105.8841
13.2177 | -105.8826
13.2214 | -105.8815
13.2252 | -105.879
13.2289 | -105.8773
13.2327 | -105.8722
13.2364 | -105.8707
13.2402 | -105.8675
13.2439 | -105.8652
13.2477 | -105.8596
13.2514 | -105.8563
13.2552 | -105.8563
13.2589 | -105.8567
13.2627 | -105.8587
13.2664 | -105.8572
13.2702 | -105.8557
13.2739 | -105.8572
13.2777 | -105.858
13.2814 | -105.8544
13.2852 | -105.8509
13.2889 | -105.8459
13.2927 | -105.8405
13.2965 | -105.8345
13.3002 | -105.8299
13.304 | -105.8166
13.3077 | -105.7992
13.3115 | -105.7804
13.3152 | -105.7629
13.319 | -105.7421
13.3227 | -105.719
13.3265 | -105.6973
13.3302 | -105.6742
13.334 | -105.6474
13.3377 | -105.6188
13.3415 | -105.5934
13.3452 | -105.5661
13.349 | -105.5395
13.3527 | -105.5142
13.3565 | -105.4897
13.3602 | -105.4663
13.364 | -105.4463
13.3677 | -105.4284
13.3715 | -105.4128
13.3752 | -105.3978
13.379 | -105.384
13.3827 | -105.3743
13.3865 | -105.3631
13.3902 | -105.3551
13.394 | -105.3494
13.3977 | -105.3447
13.4015 | -105.3433
13.4052 | -105.3429
13.409 | -105.3406
13.4127 | -105.3407
13.4165 | -105.3424
13.4202 | -105.3456
13.424 | -105.35
13.4277 | -105.355
13.4315 | -105.3582
13.4352 | -105.3591
13.439 | -105.3599
13.4427 | -105.3572
13.4465 | -105.3509
13.4502 | -105.3423
13.454 | -105.3279
13.4577 | -105.3132
13.4615 | -105.2951
13.4652 | -105.2758
13.469 | -105.2556
13.4727 | -105.2339
13.4765 | -105.2102
13.4802 | -105.1855
13.484 | -105.1578
13.4877 | -105.1314
13.4915 | -105.1038
13.4952 | -105.079
13.499 | -105.056
13.5027 | -105.0344
13.5065 | -105.0097
13.5102 | -104.987
13.514 | -104.969
13.5177 | -104.9526
13.5215 | -104.9372
13.5252 | -104.9331
13.529 | -104.924
13.5327 | -104.9175
13.5365 | -104.9126
13.5403 | -104.9105
13.544 | -104.9129
13.5478 | -104.9159
13.5515 | -104.9198
13.5553 | -104.9296
13.559 | -104.9417
13.5628 | -104.9562
13.5665 | -104.9761
13.5703 | -104.9986
13.574 | -105.0237
13.5778 | -105.0497
13.5815 | -105.0792
13.5853 | -105.1083
13.589 | -105.1365
13.5928 | -105.163
13.5965 | -105.1916
13.6003 | -105.2204
13.604 | -105.2534
13.6078 | -105.2832
13.6115 | -105.3078
13.6153 | -105.3291
13.619 | -105.3502
13.6228 | -105.3687
13.6265 | -105.3834
13.6303 | -105.3971
13.634 | -105.4077
13.6378 | -105.4151
13.6415 | -105.4181
13.6453 | -105.418
13.649 | -105.4174
13.6528 | -105.4141
13.6565 | -105.4119
13.6603 | -105.405
13.664 | -105.397
13.6678 | -105.3892
13.6715 | -105.3795
13.6753 | -105.3691
13.679 | -105.3593
13.6828 | -105.3499
13.6865 | -105.3406
13.6903 | -105.3314
13.694 | -105.323
13.6978 | -105.3166
13.7015 | -105.3149
13.7053 | -105.3145
13.709 | -105.3118
13.7128 | -105.3111
13.7165 | -105.3128
13.7203 | -105.3199
13.724 | -105.3278
13.7278 | -105.3396
13.7315 | -105.3496
13.7353 | -105.3609
13.739 | -105.3719
13.7428 | -105.3842
13.7465 | -105.3969
13.7503 | -105.4121
13.754 | -105.4262
13.7578 | -105.4383
13.7615 | -105.4494
13.7653 | -105.4549
13.769 | -105.461
13.7728 | -105.4638
13.7765 | -105.4638
13.7803 | -105.4592
13.7841 | -105.4539
13.7878 | -105.4483
13.7916 | -105.4436
13.7953 | -105.4311
13.7991 | -105.4158
13.8028 | -105.3951
13.8066 | -105.3752
13.8103 | -105.3581
13.8141 | -105.3394
13.8178 | -105.3192
13.8216 | -105.2986
13.8253 | -105.2824
13.8291 | -105.2664
13.8328 | -105.2527
13.8366 | -105.2397
13.8403 | -105.2309
13.8441 | -105.2223
13.8478 | -105.2145
13.8516 | -105.2079
13.8553 | -105.2042
13.8591 | -105.2012
13.8628 | -105.2003
13.8666 | -105.2021
13.8703 | -105.207
13.8741 | -105.2154
13.8778 | -105.2252
13.8816 | -105.2337
13.8853 | -105.2454
13.8891 | -105.2591
13.8928 | -105.2731
13.8966 | -105.2863
13.9003 | -105.302
13.9041 | -105.3181
13.9078 | -105.3371
13.9116 | -105.354
13.9153 | -105.3699
13.9191 | -105.3863
13.9228 | -105.4017
13.9266 | -105.4123
13.9303 | -105.4256
13.9341 | -105.4384
13.9378 | -105.4508
13.9416 | -105.4565
13.9453 | -105.4632
13.9491 | -105.4654
13.9528 | -105.4649
13.9566 | -105.4658
13.9603 | -105.4609
13.9641 | -105.4538
13.9678 | -105.4447
13.9716 | -105.4385
13.9753 | -105.4316
13.9791 | -105.4244
13.9828 | -105.4138
13.9866 | -105.4016
13.9903 | -105.3877
13.9941 | -105.3723
13.9978 | -105.3563
14.0016 | -105.3452
14.0053 | -105.3319
14.0091 | -105.3161
14.0128 | -105.301
14.0166 | -105.2858
14.0203 | -105.2685
14.0241 | -105.2522
14.0278 | -105.2373
14.0316 | -105.2214
14.0354 | -105.2055
14.0391 | -105.1926
14.0429 | -105.182
14.0466 | -105.1753
14.0504 | -105.1722
14.0541 | -105.1696
14.0579 | -105.1704
14.0616 | -105.1746
14.0654 | -105.1953
14.0691 | -105.2035
14.0729 | -105.213
14.0766 | -105.2243
14.0804 | -105.2399
14.0841 | -105.2584
14.0879 | -105.2775
14.0916 | -105.2994
14.0954 | -105.3218
14.0991 | -105.3434
14.1029 | -105.3628
14.1066 | -105.3852
14.1104 | -105.4093
14.1141 | -105.4347
14.1179 | -105.4561
14.1216 | -105.4736
14.1254 | -105.4927
14.1291 | -105.5125
14.1329 | -105.5315
14.1366 | -105.5506
14.1404 | -105.5702
14.1441 | -105.5927
14.1479 | -105.6092
14.1516 | -105.6265
14.1554 | -105.6455
14.1591 | -105.6616
14.1629 | -105.6755
14.1666 | -105.69
14.1704 | -105.7051
14.1741 | -105.7151
14.1779 | -105.7241
14.1816 | -105.7293
14.1854 | -105.7346
14.1891 | -105.74
14.1929 | -105.7455
14.1966 | -105.7537
14.2004 | -105.7602
14.2041 | -105.7599
14.2079 | -105.7594
14.2116 | -105.7616
14.2154 | -105.762
14.2191 | -105.7604
14.2229 | -105.7627
14.2266 | -105.7638
14.2304 | -105.7637
14.2341 | -105.7639
14.2379 | -105.7643
14.2416 | -105.7611
14.2454 | -105.758
14.2491 | -105.7545
14.2529 | -105.7513
14.2566 | -105.7478
14.2604 | -105.7485
14.2641 | -105.7473
14.2679 | -105.7446
14.2716 | -105.7429
14.2754 | -105.7439
14.2791 | -105.7447
14.2829 | -105.7455
14.2867 | -105.7447
14.2904 | -105.7446
14.2942 | -105.744
14.2979 | -105.7434
14.3017 | -105.7449
14.3054 | -105.7473
14.3092 | -105.7517
14.3129 | -105.7538
14.3167 | -105.7547
14.3204 | -105.7551
14.3242 | -105.7587
14.3279 | -105.7613
14.3317 | -105.7628
14.3354 | -105.7664
14.3392 | -105.7683
14.3429 | -105.7688
14.3467 | -105.7727
14.3504 | -105.7766
14.3542 | -105.7845
14.3579 | -105.7916
14.3617 | -105.7964
14.3654 | -105.8044
14.3692 | -105.8166
14.3729 | -105.8212
14.3767 | -105.8218
14.3804 | -105.8261
14.3842 | -105.8307
14.3879 | -105.836
14.3917 | -105.8372
14.3954 | -105.8354
14.3992 | -105.8338
14.4029 | -105.831
14.4067 | -105.8264
14.4104 | -105.8237
14.4142 | -105.8196
14.4179 | -105.8117
14.4217 | -105.8042
14.4254 | -105.7948
14.4292 | -105.7817
14.4329 | -105.7651
14.4367 | -105.7484
14.4404 | -105.7347
14.4442 | -105.7231
14.4479 | -105.7129
14.4517 | -105.7005
14.4554 | -105.6859
14.4592 | -105.6742
14.4629 | -105.6653
14.4667 | -105.6561
14.4704 | -105.6496
14.4742 | -105.6472
14.4779 | -105.6422
14.4817 | -105.6384
14.4854 | -105.6363
14.4892 | -105.6362
14.4929 | -105.6372
14.4967 | -105.6396
14.5004 | -105.6416
14.5042 | -105.6455
14.5079 | -105.651
14.5117 | -105.6554
14.5154 | -105.6543
14.5192 | -105.6533
14.5229 | -105.6523
14.5267 | -105.6518
14.5305 | -105.6503
14.5342 | -105.6499
14.538 | -105.6455
14.5417 | -105.6379
14.5455 | -105.6298
14.5492 | -105.6231
14.553 | -105.6127
14.5567 | -105.6005
14.5605 | -105.5879
14.5642 | -105.5763
14.568 | -105.5646
14.5717 | -105.552
14.5755 | -105.5413
14.5792 | -105.5342
14.583 | -105.5261
14.5867 | -105.514
14.5905 | -105.5009
14.5942 | -105.4918
14.598 | -105.4862
14.6017 | -105.4804
14.6055 | -105.4719
14.6092 | -105.4776
14.613 | -105.4759
14.6167 | -105.4729
14.6205 | -105.4709
14.6242 | -105.4726
14.628 | -105.4757
14.6317 | -105.4804
14.6355 | -105.4859
14.6392 | -105.4933
14.643 | -105.5003
14.6467 | -105.5077
14.6505 | -105.5133
14.6542 | -105.5183
14.658 | -105.5231
14.6617 | -105.5301
14.6655 | -105.5367
14.6692 | -105.5393
14.673 | -105.5418
14.6767 | -105.5468
14.6805 | -105.5517
14.6842 | -105.5538
14.688 | -105.5594
14.6917 | -105.5674
14.6955 | -105.5745
14.6992 | -105.5824
14.703 | -105.5904
14.7067 | -105.5985
14.7105 | -105.6079
14.7142 | -105.6158
14.718 | -105.6236
14.7217 | -105.6328
14.7255 | -105.645
14.7292 | -105.6561
14.733 | -105.6679
14.7367 | -105.6784
14.7405 | -105.6887
14.7442 | -105.6983
14.748 | -105.7061
14.7517 | -105.7116
14.7555 | -105.7195
14.7592 | -105.7232
14.763 | -105.7236
14.7667 | -105.7235
14.7705 | -105.7252
14.7743 | -105.7261
14.778 | -105.7235
14.7818 | -105.7151
14.7855 | -105.7059
14.7893 | -105.6962
14.793 | -105.6851
14.7968 | -105.6753
14.8005 | -105.6618
14.8043 | -105.642
14.808 | -105.6202
14.8118 | -105.5963
14.8155 | -105.5719
14.8193 | -105.5432
14.823 | -105.514
14.8268 | -105.4849
14.8305 | -105.4593
14.8343 | -105.4341
14.838 | -105.4077
14.8418 | -105.3805
14.8455 | -105.3538
14.8493 | -105.3319
14.853 | -105.3093
14.8568 | -105.2925
14.8605 | -105.2811
14.8643 | -105.2727
14.868 | -105.266
14.8718 | -105.2594
14.8755 | -105.2569
14.8793 | -105.2586
14.883 | -105.2605
14.8868 | -105.2623
14.8905 | -105.2664
14.8943 | -105.2745
14.898 | -105.2851
14.9018 | -105.2957
14.9055 | -105.3037
14.9093 | -105.3123
14.913 | -105.3205
14.9168 | -105.3287
14.9205 | -105.3368
14.9243 | -105.3443
14.928 | -105.3512
14.9318 | -105.3571
14.9355 | -105.3637
14.9393 | -105.368
14.943 | -105.3708
14.9468 | -105.3705
14.9505 | -105.3692
14.9543 | -105.3663
14.958 | -105.3618
14.9618 | -105.3563
14.9655 | -105.3496
14.9693 | -105.3398
14.973 | -105.3303
14.9768 | -105.3223
14.9805 | -105.31
14.9843 | -105.2929
14.988 | -105.2768
14.9918 | -105.2643
14.9955 | -105.2548
14.9993 | -105.2423
15.003 | -105.2297
15.0068 | -105.2174
15.0105 | -105.2035
15.0143 | -105.1931
15.0181 | -105.1856
15.0218 | -105.1764
15.0256 | -105.1652
15.0293 | -105.156
15.0331 | -105.149
15.0368 | -105.1415
15.0406 | -105.1378
15.0443 | -105.1347
15.0481 | -105.1308
15.0518 | -105.1264
15.0556 | -105.1225
15.0593 | -105.1195
15.0631 | -105.1164
15.0668 | -105.115
15.0706 | -105.1113
15.0743 | -105.104
15.0781 | -105.0964
15.0818 | -105.0889
15.0856 | -105.0831
15.0893 | -105.0744
15.0931 | -105.0629
15.0968 | -105.0513
15.1006 | -105.038
15.1043 | -105.0219
15.1081 | -105.0022
15.1118 | -104.9833
15.1156 | -104.9652
15.1193 | -104.949
15.1231 | -104.9338
15.1268 | -104.9175
15.1306 | -104.9006
15.1343 | -104.884
15.1381 | -104.868
15.1418 | -104.8543
15.1456 | -104.8409
15.1493 | -104.8393
15.1531 | -104.8315
15.1568 | -104.8252
15.1606 | -104.822
15.1643 | -104.8217
15.1681 | -104.8232
15.1718 | -104.8256
15.1756 | -104.8306
15.1793 | -104.8413
15.1831 | -104.8531
15.1868 | -104.8649
15.1906 | -104.8774
15.1943 | -104.8927
15.1981 | -104.9138
15.2018 | -104.9347
15.2056 | -104.9561
15.2093 | -104.9792
15.2131 | -105.0013
15.2168 | -105.0258
15.2206 | -105.0519
15.2243 | -105.0779
15.2281 | -105.1036
15.2318 | -105.1317
15.2356 | -105.1588
15.2393 | -105.1896
15.2431 | -105.2224
15.2468 | -105.2514
15.2506 | -105.2761
15.2543 | -105.2993
15.2581 | -105.3223
15.2619 | -105.3437
15.2656 | -105.3657
15.2694 | -105.3893
15.2731 | -105.413
15.2769 | -105.4332
15.2806 | -105.4508
15.2844 | -105.4679
15.2881 | -105.4836
15.2919 | -105.499
15.2956 | -105.5168
15.2994 | -105.5336
15.3031 | -105.5458
15.3069 | -105.558
15.3106 | -105.571
15.3144 | -105.5843
15.3181 | -105.596
15.3219 | -105.6057
15.3256 | -105.6135
15.3294 | -105.6161
15.3331 | -105.6156
15.3369 | -105.6164
15.3406 | -105.6161
15.3444 | -105.6129
15.3481 | -105.6082
15.3519 | -105.6005
15.3556 | -105.592
15.3594 | -105.5796
15.3631 | -105.5679
15.3669 | -105.5575
15.3706 | -105.5466
15.3744 | -105.5316
15.3781 | -105.5176
15.3819 | -105.5034
15.3856 | -105.4876
15.3894 | -105.4737
15.3931 | -105.4617
15.3969 | -105.4482
15.4006 | -105.4312
15.4044 | -105.4189
15.4081 | -105.4075
15.4119 | -105.3998
15.4156 | -105.393
15.4194 | -105.3866
15.4231 | -105.3823
15.4269 | -105.3804
15.4306 | -105.3795
15.4344 | -105.3819
15.4381 | -105.3863
15.4419 | -105.3916
15.4456 | -105.4014
15.4494 | -105.4131
15.4531 | -105.4207
15.4569 | -105.4294
15.4606 | -105.4392
15.4644 | -105.4519
15.4681 | -105.4638
15.4719 | -105.4766
15.4756 | -105.4885
15.4794 | -105.497
15.4831 | -105.506
15.4869 | -105.5116
15.4906 | -105.5165
15.4944 | -105.5182
15.4981 | -105.5206
15.5019 | -105.5228
15.5056 | -105.5192
15.5094 | -105.5122
15.5132 | -105.5018
15.5169 | -105.4859
15.5207 | -105.4673
15.5244 | -105.4525
15.5282 | -105.4358
15.5319 | -105.4129
15.5357 | -105.382
15.5394 | -105.3493
15.5432 | -105.3178
15.5469 | -105.2866
15.5507 | -105.2515
15.5544 | -105.2117
15.5582 | -105.1694
15.5619 | -105.1258
15.5657 | -105.0825
15.5694 | -105.0376
15.5732 | -104.9915
15.5769 | -104.9457
15.5807 | -104.9007
15.5844 | -104.8537
15.5882 | -104.8045
15.5919 | -104.7579
15.5957 | -104.7139
15.5994 | -104.6727
15.6032 | -104.6323
15.6069 | -104.5951
15.6107 | -104.5593
15.6144 | -104.5268
15.6182 | -104.4973
15.6219 | -104.4686
15.6257 | -104.4432
15.6294 | -104.4234
15.6332 | -104.4055
15.6369 | -104.39
15.6407 | -104.3747
15.6444 | -104.364
15.6482 | -104.3537
15.6519 | -104.3457
15.6557 | -104.3404
15.6594 | -104.3375
15.6632 | -104.3377
15.6669 | -104.3373
15.6707 | -104.3388
15.6744 | -104.3434
15.6782 | -104.3468
15.6819 | -104.3486
15.6857 | -104.3509
15.6894 | -104.3522
15.6932 | -104.3542
15.6969 | -104.356
15.7007 | -104.3578
15.7044 | -104.3595
15.7082 | -104.3617
15.7119 | -104.3649
15.7157 | -104.368
15.7194 | -104.3731
15.7232 | -104.3751
15.7269 | -104.3763
15.7307 | -104.3799
15.7344 | -104.3852
15.7382 | -104.3908
15.7419 | -104.3984
15.7457 | -104.408
15.7494 | -104.4157
15.7532 | -104.4258
15.757 | -104.4389
15.7607 | -104.4568
15.7645 | -104.4775
15.7682 | -104.4959
15.772 | -104.5145
15.7757 | -104.537
15.7795 | -104.5624
15.7832 | -104.5869
15.787 | -104.6115
15.7907 | -104.6375
15.7945 | -104.6625
15.7982 | -104.6879
15.802 | -104.7109
15.8057 | -104.7355
15.8095 | -104.7589
15.8132 | -104.7812
15.817 | -104.7988
15.8207 | -104.8178
15.8245 | -104.8367
15.8282 | -104.854
15.832 | -104.8692
15.8357 | -104.8849
15.8395 | -104.8981
15.8432 | -104.9099
15.847 | -104.9249
15.8507 | -104.9358
15.8545 | -104.9436
15.8582 | -104.9511
15.862 | -104.9612
15.8657 | -104.9736
15.8695 | -104.984
15.8732 | -104.9951
15.877 | -105.0054
15.8807 | -105.0148
15.8845 | -105.0239
15.8882 | -105.0319
15.892 | -105.0402
15.8957 | -105.0483
15.8995 | -105.056
15.9032 | -105.0675
15.907 | -105.0792
15.9107 | -105.09
15.9145 | -105.1007
15.9182 | -105.1081
15.922 | -105.1162
15.9257 | -105.1268
15.9295 | -105.1362
15.9332 | -105.1451
15.937 | -105.1517
15.9407 | -105.1559
15.9445 | -105.1549
15.9482 | -105.1553
15.952 | -105.1555
15.9557 | -105.1531
15.9595 | -105.1505
15.9632 | -105.146
15.967 | -105.1398
15.9707 | -105.1301
15.9745 | -105.1166
15.9782 | -105.1039
15.982 | -105.087
15.9857 | -105.0704
15.9895 | -105.0553
15.9932 | -105.0372
15.997 | -105.0163
16.0008 | -104.9965
16.0045 | -104.9791
16.0083 | -104.9623
16.012 | -104.9434
16.0158 | -104.9268
16.0195 | -104.913
16.0233 | -104.8977
16.027 | -104.8821
16.0308 | -104.8693
16.0345 | -104.8538
16.0383 | -104.8423
16.042 | -104.8351
16.0458 | -104.8307
16.0495 | -104.8242
16.0533 | -104.8169
16.057 | -104.8118
16.0608 | -104.8066
16.0645 | -104.7976
16.0683 | -104.7929
16.072 | -104.7912
16.0758 | -104.7881
16.0795 | -104.784
16.0833 | -104.7821
16.087 | -104.78
16.0908 | -104.7754
16.0945 | -104.7718
16.0983 | -104.7683
16.102 | -104.7654
16.1058 | -104.7639
16.1095 | -104.7621
16.1133 | -104.7607
16.117 | -104.759
16.1208 | -104.7573
16.1245 | -104.7555
16.1283 | -104.755
16.132 | -104.754
16.1358 | -104.755
16.1395 | -104.7567
16.1433 | -104.7603
16.147 | -104.769
16.1508 | -104.7773
16.1545 | -104.7842
16.1583 | -104.7941
16.162 | -104.8062
16.1658 | -104.8197
16.1695 | -104.8353
16.1733 | -104.8538
16.177 | -104.8725
16.1808 | -104.8932
16.1845 | -104.9154
16.1883 | -104.9373
16.192 | -104.9597
16.1958 | -104.9835
16.1995 | -105.0088
16.2033 | -105.0364
16.207 | -105.0623
16.2108 | -105.0892
16.2145 | -105.1129
16.2183 | -105.1337
16.222 | -105.1569
16.2258 | -105.1834
16.2295 | -105.2079
16.2333 | -105.2251
16.237 | -105.2423
16.2408 | -105.2561
16.2446 | -105.2731
16.2483 | -105.2894
16.2521 | -105.3005
16.2558 | -105.3139
16.2596 | -105.3261
16.2633 | -105.3337
16.2671 | -105.3421
16.2708 | -105.3473
16.2746 | -105.347
16.2783 | -105.3461
16.2821 | -105.3463
16.2858 | -105.3435
16.2896 | -105.3387
16.2933 | -105.334
16.2971 | -105.3282
16.3008 | -105.3156
16.3046 | -105.3002
16.3083 | -105.2854
16.3121 | -105.2705
16.3158 | -105.2524
16.3196 | -105.2344
16.3233 | -105.2176
16.3271 | -105.2019
16.3308 | -105.1865
16.3346 | -105.1701
16.3383 | -105.153
16.3421 | -105.1348
16.3458 | -105.1177
16.3496 | -105.1046
16.3533 | -105.0956
16.3571 | -105.0873
16.3608 | -105.0792
16.3646 | -105.0723
16.3683 | -105.0665
16.3721 | -105.0645
16.3758 | -105.0652
16.3796 | -105.0671
16.3833 | -105.072
16.3871 | -105.08
16.3908 | -105.0891
16.3946 | -105.0985
16.3983 | -105.1116
16.4021 | -105.127
16.4058 | -105.1449
16.4096 | -105.163
16.4133 | -105.1819
16.4171 | -105.2014
16.4208 | -105.2217
16.4246 | -105.2432
16.4283 | -105.2637
16.4321 | -105.2829
16.4358 | -105.3002
16.4396 | -105.3165
16.4433 | -105.3315
16.4471 | -105.3494
16.4508 | -105.3641
16.4546 | -105.3722
16.4583 | -105.38
16.4621 | -105.3873
16.4658 | -105.3928
16.4696 | -105.3948
16.4733 | -105.395
16.4771 | -105.396
16.4808 | -105.3913
16.4846 | -105.3827
16.4883 | -105.374
16.4921 | -105.3633
16.4959 | -105.3507
16.4996 | -105.3366
16.5034 | -105.3196
16.5071 | -105.3013
16.5109 | -105.2801
16.5146 | -105.2563
16.5184 | -105.2316
16.5221 | -105.2103
16.5259 | -105.1912
16.5296 | -105.1734
16.5334 | -105.1574
16.5371 | -105.1415
16.5409 | -105.1219
16.5446 | -105.1043
16.5484 | -105.0874
16.5521 | -105.0713
16.5559 | -105.0562
16.5596 | -105.0445
16.5634 | -105.0329
16.5671 | -105.0245
16.5709 | -105.0198
16.5746 | -105.0166
16.5784 | -105.0177
16.5821 | -105.0177
16.5859 | -105.0164
16.5896 | -105.0164
16.5934 | -105.0227
16.5971 | -105.0295
16.6009 | -105.0391
16.6046 | -105.0504
16.6084 | -105.0619
16.6121 | -105.0747
16.6159 | -105.0892
16.6196 | -105.1026
16.6234 | -105.1171
16.6271 | -105.1338
16.6309 | -105.1535
16.6346 | -105.1708
16.6384 | -105.185
16.6421 | -105.1979
16.6459 | -105.2045
16.6496 | -105.2167
16.6534 | -105.2292
16.6571 | -105.2372
16.6609 | -105.2442
16.6646 | -105.2503
16.6684 | -105.2547
16.6721 | -105.2571
16.6759 | -105.2538
16.6796 | -105.2473
16.6834 | -105.2372
16.6871 | -105.2267
16.6909 | -105.2157
16.6946 | -105.1995
16.6984 | -105.1794
16.7021 | -105.1536
16.7059 | -105.1281
16.7096 | -105.105
16.7134 | -105.0776
16.7171 | -105.0496
16.7209 | -105.0206
16.7246 | -104.9917
16.7284 | -104.9641
16.7321 | -104.935
16.7359 | -104.9059
16.7397 | -104.8763
16.7434 | -104.8497
16.7472 | -104.8276
16.7509 | -104.8084
16.7547 | -104.7915
16.7584 | -104.776
16.7622 | -104.7595
16.7659 | -104.7463
16.7697 | -104.7375
16.7734 | -104.748
16.7772 | -104.7448
16.7809 | -104.7477
16.7847 | -104.7494
16.7884 | -104.7556
16.7922 | -104.7656
16.7959 | -104.7761
16.7997 | -104.7844
16.8034 | -104.7971
16.8072 | -104.8115
16.8109 | -104.8252
16.8147 | -104.8401
16.8184 | -104.8566
16.8222 | -104.8741
16.8259 | -104.8921
16.8297 | -104.9096
16.8334 | -104.9236
16.8372 | -104.937
16.8409 | -104.9477
16.8447 | -104.9595
16.8484 | -104.9723
16.8522 | -104.9828
16.8559 | -104.9902
16.8597 | -104.9976
16.8634 | -105.0018
16.8672 | -105.0047
16.8709 | -105.007
16.8747 | -105.0075
16.8784 | -105.0084
16.8822 | -105.0077
16.8859 | -105.0091
16.8897 | -105.0014
16.8934 | -104.992
16.8972 | -104.9833
16.9009 | -104.9712
16.9047 | -104.9611
16.9084 | -104.9502
16.9122 | -104.9391
16.9159 | -104.9293
16.9197 | -104.9176
16.9234 | -104.9048
16.9272 | -104.8923
16.9309 | -104.8794
16.9347 | -104.8635
16.9384 | -104.8476
16.9422 | -104.8333
16.9459 | -104.8149
16.9497 | -104.7939
16.9534 | -104.7722
16.9572 | -104.746
16.9609 | -104.7179
16.9647 | -104.6918
16.9684 | -104.6629
16.9722 | -104.6329
16.9759 | -104.6024
16.9797 | -104.5707
16.9835 | -104.5335
16.9872 | -104.4931
16.991 | -104.4442
16.9947 | -104.3947
16.9985 | -104.3478
17.0022 | -104.303
17.006 | -104.252
17.0097 | -104.1983
17.0135 | -104.1389
17.0172 | -104.0792
17.021 | -104.023
17.0247 | -103.9655
17.0285 | -103.9066
17.0322 | -103.8442
17.036 | -103.7822
17.0397 | -103.722
17.0435 | -103.6638
17.0472 | -103.6058
17.051 | -103.5491
17.0547 | -103.494
17.0585 | -103.4397
17.0622 | -103.3876
17.066 | -103.3361
17.0697 | -103.2869
17.0735 | -103.2396
17.0772 | -103.1967
17.081 | -103.1582
17.0847 | -103.122
17.0885 | -103.0889
17.0922 | -103.0585
17.096 | -103.0316
17.0997 | -103.0097
17.1035 | -102.9899
17.1072 | -102.9731
17.111 | -102.9589
17.1147 | -102.9454
17.1185 | -102.9352
17.1222 | -102.9274
17.126 | -102.9206
17.1297 | -102.9123
17.1335 | -102.9052
17.1372 | -102.902
17.141 | -102.9008
17.1447 | -102.8992
17.1485 | -102.8965
17.1522 | -102.8948
17.156 | -102.8931
17.1597 | -102.8918
17.1635 | -102.8908
17.1672 | -102.89
17.171 | -102.8899
17.1747 | -102.8908
17.1785 | -102.8948
17.1822 | -102.8984
17.186 | -102.9025
17.1897 | -102.9103
17.1935 | -102.9181
17.1972 | -102.9296
17.201 | -102.9455
17.2047 | -102.9614
17.2085 | -102.9747
17.2122 | -102.9928
17.216 | -103.014
17.2197 | -103.0368
17.2235 | -103.0619
17.2272 | -103.0912
17.231 | -103.1222
17.2348 | -103.1587
17.2385 | -103.1928
17.2423 | -103.2204
17.246 | -103.2499
17.2498 | -103.2845
17.2535 | -103.3179
17.2573 | -103.356
17.261 | -103.3885
17.2648 | -103.4139
17.2685 | -103.4367
17.2723 | -103.4585
17.276 | -103.4821
17.2798 | -103.5007
17.2835 | -103.5171
17.2873 | -103.5322
17.291 | -103.5458
17.2948 | -103.5557
17.2985 | -103.5629
17.3023 | -103.5653
17.306 | -103.5648
17.3098 | -103.5634
17.3135 | -103.536
17.3173 | -103.5268
17.321 | -103.5139
17.3248 | -103.4981
17.3285 | -103.4787
17.3323 | -103.4562
17.336 | -103.432
17.3398 | -103.4066
17.3435 | -103.3811
17.3473 | -103.3539
17.351 | -103.3244
17.3548 | -103.2963
17.3585 | -103.2675
17.3623 | -103.2404
17.366 | -103.2154
17.3698 | -103.1901
17.3735 | -103.1667
17.3773 | -103.146
17.381 | -103.1256
17.3848 | -103.1027
17.3885 | -103.0815
17.3923 | -103.0605
17.396 | -103.0424
17.3998 | -103.0264
17.4035 | -103.0121
17.4073 | -103.0001
17.411 | -102.9891
17.4148 | -102.9778
17.4185 | -102.9645
17.4223 | -102.9533
17.426 | -102.9455
17.4298 | -102.9392
17.4335 | -102.9325
17.4373 | -102.9259
17.441 | -102.9193
17.4448 | -102.9129
17.4485 | -102.9062
17.4523 | -102.9015
17.456 | -102.8998
17.4598 | -102.8961
17.4635 | -102.8918
17.4673 | -102.8869
17.471 | -102.8837
17.4748 | -102.8812
17.4785 | -102.8801
17.4823 | -102.8803
17.4861 | -102.8817
17.4898 | -102.8866
17.4936 | -102.892
17.4973 | -102.8969
17.5011 | -102.9041
17.5048 | -102.9121
17.5086 | -102.9231
17.5123 | -102.9322
17.5161 | -102.9438
17.5198 | -102.9602
17.5236 | -102.9763
17.5273 | -102.9954
17.5311 | -103.0141
17.5348 | -103.034
17.5386 | -103.054
17.5423 | -103.079
17.5461 | -103.1056
17.5498 | -103.1364
17.5536 | -103.1682
17.5573 | -103.2009
17.5611 | -103.2353
17.5648 | -103.2709
17.5686 | -103.3105
17.5723 | -103.3509
17.5761 | -103.3873
17.5798 | -103.4234
17.5836 | -103.462
17.5873 | -103.5
17.5911 | -103.538
17.5948 | -103.5744
17.5986 | -103.6091
17.6023 | -103.6454
17.6061 | -103.6768
17.6098 | -103.7057
17.6136 | -103.7349
17.6173 | -103.7598
17.6211 | -103.7851
17.6248 | -103.808
17.6286 | -103.8282
17.6323 | -103.8467
17.6361 | -103.865
17.6398 | -103.8808
17.6436 | -103.8966
17.6473 | -103.9083
17.6511 | -103.9193
17.6548 | -103.9303
17.6586 | -103.9356
17.6623 | -103.9392
17.6661 | -103.9428
17.6698 | -103.9492
17.6736 | -103.9539
17.6773 | -103.954
17.6811 | -103.9499
17.6848 | -103.9437
17.6886 | -103.9392
17.6923 | -103.9367
17.6961 | -103.9331
17.6998 | -103.9307
17.7036 | -103.9284
17.7073 | -103.923
17.7111 | -103.915
17.7148 | -103.909
17.7186 | -103.9075
17.7223 | -103.9043
17.7261 | -103.9002
17.7299 | -103.8956
17.7336 | -103.8905
17.7374 | -103.8879
17.7411 | -103.8877
17.7449 | -103.8865
17.7486 | -103.8862
17.7524 | -103.888
17.7561 | -103.8897
17.7599 | -103.8903
17.7636 | -103.8981
17.7674 | -103.9078
17.7711 | -103.9141
17.7749 | -103.921
17.7786 | -103.9296
17.7824 | -103.9403
17.7861 | -103.9528
17.7899 | -103.969
17.7936 | -103.9845
17.7974 | -104.0047
17.8011 | -104.025
17.8049 | -104.0466
17.8086 | -104.0726
17.8124 | -104.1022
17.8161 | -104.1344
17.8199 | -104.1648
17.8236 | -104.1967
17.8274 | -104.232
17.8311 | -104.2671
17.8349 | -104.2985
17.8386 | -104.3302
17.8424 | -104.3616
17.8461 | -104.394
17.8499 | -104.4257
17.8536 | -104.4403
17.8574 | -104.4686
17.8611 | -104.4953
17.8649 | -104.5179
17.8686 | -104.5361
17.8724 | -104.5525
17.8761 | -104.565
17.8799 | -104.5753
17.8836 | -104.5856
17.8874 | -104.5919
17.8911 | -104.5932
17.8949 | -104.5914
17.8986 | -104.5868
17.9024 | -104.579
17.9061 | -104.567
17.9099 | -104.5545
17.9136 | -104.5433
17.9174 | -104.5317
17.9211 | -104.517
17.9249 | -104.4972
17.9286 | -104.4768
17.9324 | -104.4581
17.9361 | -104.4434
17.9399 | -104.4292
17.9436 | -104.4097
17.9474 | -104.3858
17.9511 | -104.3617
17.9549 | -104.3385
17.9586 | -104.318
17.9624 | -104.2971
17.9661 | -104.281
17.9699 | -104.264
17.9737 | -104.2449
17.9774 | -104.2277
17.9812 | -104.2126
17.9849 | -104.2008
17.9887 | -104.1892
17.9924 | -104.1779
17.9962 | -104.166
17.9999 | -104.1565
18.0037 | -104.15
18.0074 | -104.145
18.0112 | -104.1419
18.0149 | -104.1421
18.0187 | -104.1406
18.0224 | -104.1373
18.0262 | -104.1361
18.0299 | -104.1378
18.0337 | -104.1404
18.0374 | -104.1445
18.0412 | -104.1497
18.0449 | -104.1541
18.0487 | -104.1592
18.0524 | -104.1615
18.0562 | -104.1693
18.0599 | -104.1774
18.0637 | -104.1834
18.0674 | -104.1891
18.0712 | -104.1937
18.0749 | -104.2015
18.0787 | -104.2124
18.0824 | -104.225
18.0862 | -104.2386
18.0899 | -104.2493
18.0937 | -104.2552
18.0974 | -104.262
18.1012 | -104.2694
18.1049 | -104.2771
18.1087 | -104.2869
18.1124 | -104.2932
18.1162 | -104.2928
18.1199 | -104.291
18.1237 | -104.2936
18.1274 | -104.2958
18.1312 | -104.2973
18.1349 | -104.2958
18.1387 | -104.2884
18.1424 | -104.2756
18.1462 | -104.2645
18.1499 | -104.2548
18.1537 | -104.2449
18.1574 | -104.2315
18.1612 | -104.2159
18.1649 | -104.1972
18.1687 | -104.172
18.1724 | -104.148
18.1762 | -104.1226
18.1799 | -104.0983
18.1837 | -104.0753
18.1874 | -104.0486
18.1912 | -104.0187
18.1949 | -103.9909
18.1987 | -103.9639
18.2024 | -103.9371
18.2062 | -103.9116
18.2099 | -103.8867
18.2137 | -103.8615
18.2174 | -103.8389
18.2212 | -103.8196
18.225 | -103.8002
18.2287 | -103.7842
18.2325 | -103.7696
18.2362 | -103.756
18.24 | -103.7444
18.2437 | -103.7336
18.2475 | -103.7265
18.2512 | -103.7249
18.255 | -103.7255
18.2587 | -103.7265
18.2625 | -103.7277
18.2662 | -103.7324
18.27 | -103.7412
18.2737 | -103.7524
18.2775 | -103.7674
18.2812 | -103.7829
18.285 | -103.7982
18.2887 | -103.8141
18.2925 | -103.8302
18.2962 | -103.8476
18.3 | -103.8657
18.3037 | -103.8828
18.3075 | -103.8958
18.3112 | -103.9071
18.315 | -103.9142
18.3187 | -103.9212
18.3225 | -103.9295
18.3262 | -103.9389
18.33 | -103.9424
18.3337 | -103.9384
18.3375 | -103.929
18.3412 | -103.918
18.345 | -103.9098
18.3487 | -103.9042
18.3525 | -103.8974
18.3562 | -103.8852
18.36 | -103.8685
18.3637 | -103.8505
18.3675 | -103.8331
18.3712 | -103.8149
18.375 | -103.7956
18.3787 | -103.7776
18.3825 | -103.7524
18.3862 | -103.7247
18.39 | -103.6992
18.3937 | -103.676
18.3975 | -103.6548
18.4012 | -103.6336
18.405 | -103.6151
18.4087 | -103.5942
18.4125 | -103.5712
18.4162 | -103.5491
18.42 | -103.5299
18.4237 | -103.5113
18.4275 | -103.4901
18.4312 | -103.4718
18.435 | -103.4552
18.4387 | -103.4412
18.4425 | -103.4277
18.4462 | -103.4135
18.45 | -103.401
18.4537 | -103.3913
18.4575 | -103.3808
18.4612 | -103.371
18.465 | -103.363
18.4688 | -103.3574
18.4725 | -103.3526
18.4763 | -103.347
18.48 | -103.3392
18.4838 | -103.3322
18.4875 | -103.3267
18.4913 | -103.3223
18.495 | -103.3189
18.4988 | -103.3144
18.5025 | -103.3073
18.5063 | -103.2972
18.51 | -103.2885
18.5138 | -103.2808
18.5175 | -103.2728
18.5213 | -103.2613
18.525 | -103.2474
18.5288 | -103.2313
18.5325 | -103.213
18.5363 | -103.198
18.54 | -103.183
18.5438 | -103.1683
18.5475 | -103.1524
18.5513 | -103.1316
18.555 | -103.1114
18.5588 | -103.0929
18.5625 | -103.0742
18.5663 | -103.0579
18.57 | -103.0431
18.5738 | -103.0283
18.5775 | -103.0101
18.5813 | -102.9919
18.585 | -102.9736
18.5888 | -102.9536
18.5925 | -102.9365
18.5963 | -102.9212
18.6 | -102.9064
18.6038 | -102.8923
18.6075 | -102.8761
18.6113 | -102.8603
18.615 | -102.8477
18.6188 | -102.8363
18.6225 | -102.8246
18.6263 | -102.8156
18.63 | -102.8093
18.6338 | -102.8054
18.6375 | -102.8065
18.6413 | -102.8084
18.645 | -102.8132
18.6488 | -102.8203
18.6525 | -102.8288
18.6563 | -102.8397
18.66 | -102.8549
18.6638 | -102.8723
18.6675 | -102.8926
18.6713 | -102.9153
18.675 | -102.9401
18.6788 | -102.966
18.6825 | -102.9941
18.6863 | -103.0263
18.69 | -103.0606
18.6938 | -103.0995
18.6975 | -103.1389
18.7013 | -103.176
18.705 | -103.2125
18.7088 | -103.2495
18.7126 | -103.2891
18.7163 | -103.3296
18.7201 | -103.3693
18.7238 | -103.4086
18.7276 | -103.451
18.7313 | -103.4925
18.7351 | -103.5296
18.7388 | -103.5639
18.7426 | -103.5994
18.7463 | -103.6341
18.7501 | -103.6691
18.7538 | -103.7038
18.7576 | -103.7353
18.7613 | -103.7654
18.7651 | -103.7965
18.7688 | -103.8264
18.7726 | -103.8559
18.7763 | -103.8809
18.7801 | -103.8994
18.7838 | -103.9177
18.7876 | -103.9382
18.7913 | -103.9585
18.7951 | -103.9765
18.7988 | -103.9907
18.8026 | -103.998
18.8063 | -104.002
18.8101 | -104.0046
18.8138 | -104.0077
18.8176 | -104.012
18.8213 | -104.011
18.8251 | -104.0072
18.8288 | -104.0019
18.8326 | -103.9976
18.8363 | -103.9931
18.8401 | -103.9851
18.8438 | -103.9773
18.8476 | -103.965
18.8513 | -103.9508
18.8551 | -103.938
18.8588 | -103.9238
18.8626 | -103.9108
18.8663 | -103.8914
18.8701 | -103.8661
18.8738 | -103.8445
18.8776 | -103.8241
18.8813 | -103.8016
18.8851 | -103.7812
18.8888 | -103.7611
18.8926 | -103.7328
18.8963 | -103.6967
18.9001 | -103.6602
18.9038 | -103.6237
18.9076 | -103.5883
18.9113 | -103.5522
18.9151 | -103.5112
18.9188 | -103.4665
18.9226 | -103.4209
18.9263 | -103.3739
18.9301 | -103.327
18.9338 | -103.2791
18.9376 | -103.2159
18.9413 | -103.163
18.9451 | -103.1108
18.9488 | -103.0587
18.9526 | -103.0039
18.9564 | -102.9422
18.9601 | -102.8797
18.9639 | -102.8186
18.9676 | -102.7611
18.9714 | -102.7073
18.9751 | -102.6533
18.9789 | -102.6008
18.9826 | -102.5474
18.9864 | -102.4973
18.9901 | -102.4495
18.9939 | -102.4036
18.9976 | -102.3591
19.0014 | -102.3173
19.0051 | -102.2777
19.0089 | -102.2399
19.0126 | -102.2057
19.0164 | -102.1747
19.0201 | -102.1443
19.0239 | -102.1151
19.0276 | -102.0881
19.0314 | -102.0641
19.0351 | -102.0422
19.0389 | -102.0254
19.0426 | -102.0121
19.0464 | -102.0007
19.0501 | -101.9912
19.0539 | -101.9816
19.0576 | -101.9737
19.0614 | -101.9695
19.0651 | -101.967
19.0689 | -101.9675
19.0726 | -101.9698
19.0764 | -101.9746
19.0801 | -101.9768
19.0839 | -101.9714
19.0876 | -101.9658
19.0914 | -101.9659
19.0951 | -101.97
19.0989 | -101.9725
19.1026 | -101.9713
19.1064 | -101.9696
19.1101 | -101.9666
19.1139 | -101.9652
19.1176 | -101.9644
19.1214 | -101.9618
19.1251 | -101.9537
19.1289 | -101.9424
19.1326 | -101.9301
19.1364 | -101.9199
19.1401 | -101.908
19.1439 | -101.8909
19.1476 | -101.8699
19.1514 | -101.8461
19.1551 | -101.8198
19.1589 | -101.7944
19.1626 | -101.7706
19.1664 | -101.7468
19.1701 | -101.7195
19.1739 | -101.6875
19.1776 | -101.6548
19.1814 | -101.6219
19.1851 | -101.5902
19.1889 | -101.5596
19.1926 | -101.5293
19.1964 | -101.4996
19.2002 | -101.4711
19.2039 | -101.4443
19.2077 | -101.4191
19.2114 | -101.3966
19.2152 | -101.3754
19.2189 | -101.3549
19.2227 | -101.3363
19.2264 | -101.3212
19.2302 | -101.3111
19.2339 | -101.3047
19.2377 | -101.3003
19.2414 | -101.2963
19.2452 | -101.2919
19.2489 | -101.2855
19.2527 | -101.2801
19.2564 | -101.2776
19.2602 | -101.2762
19.2639 | -101.2774
19.2677 | -101.2821
19.2714 | -101.2853
19.2752 | -101.2886
19.2789 | -101.2946
19.2827 | -101.3028
19.2864 | -101.3089
19.2902 | -101.3149
19.2939 | -101.3201
19.2977 | -101.3216
19.3014 | -101.3202
19.3052 | -101.3177
19.3089 | -101.3129
19.3127 | -101.3102
19.3164 | -101.3066
19.3202 | -101.2956
19.3239 | -101.2766
19.3277 | -101.2555
19.3314 | -101.234
19.3352 | -101.2129
19.3389 | -101.1927
19.3427 | -101.1665
19.3464 | -101.1321
19.3502 | -101.0942
19.3539 | -101.0601
19.3577 | -101.0334
19.3614 | -101.0066
19.3652 | -100.9725
19.3689 | -100.9334
19.3727 | -100.894
19.3764 | -100.8554
19.3802 | -100.8159
19.3839 | -100.778
19.3877 | -100.7376
19.3914 | -100.6943
19.3952 | -100.6527
19.3989 | -100.6141
19.4027 | -100.578
19.4064 | -100.5408
19.4102 | -100.5044
19.4139 | -100.4699
19.4177 | -100.4358
19.4214 | -100.4038
19.4252 | -100.3753
19.4289 | -100.3495
19.4327 | -100.3264
19.4364 | -100.3057
19.4402 | -100.287
19.444 | -100.2728
19.4477 | -100.2632
19.4515 | -100.2568
19.4552 | -100.2548
19.459 | -100.2575
19.4627 | -100.2619
19.4665 | -100.27
19.4702 | -100.2834
19.474 | -100.3019
19.4777 | -100.3535
19.4815 | -100.3809
19.4852 | -100.4134
19.489 | -100.45
19.4927 | -100.4907
19.4965 | -100.5359
19.5002 | -100.5853
19.504 | -100.6381
19.5077 | -100.6926
19.5115 | -100.7498
19.5152 | -100.8093
19.519 | -100.8708
19.5227 | -100.9364
19.5265 | -101.0084
19.5302 | -101.0811
19.534 | -101.1524
19.5377 | -101.2224
19.5415 | -101.2902
19.5452 | -101.3572
19.549 | -101.4262
19.5527 | -101.496
19.5565 | -101.5641
19.5602 | -101.6263
19.564 | -101.6789
19.5677 | -101.7257
19.5715 | -101.777
19.5752 | -101.8318
19.579 | -101.8851
19.5827 | -101.9291
19.5865 | -101.9619
19.5902 | -101.9815
19.594 | -101.999
19.5977 | -102.0224
19.6015 | -102.0476
19.6052 | -102.0657
19.609 | -102.0733
19.6127 | -102.0703
19.6165 | -102.0562
19.6202 | -102.0424
19.624 | -102.0318
19.6277 | -102.0147
19.6315 | -101.9862
19.6352 | -101.945
19.639 | -101.895
19.6427 | -101.8445
19.6465 | -101.7961
19.6502 | -101.7487
19.654 | -101.6942
19.6577 | -101.6297
19.6615 | -101.5575
19.6652 | -101.487
19.669 | -101.4246
19.6727 | -101.3662
19.6765 | -101.3091
19.6802 | -101.2505
19.684 | -101.1923
19.6877 | -101.1352
19.6915 | -101.0808
19.6953 | -101.0311
19.699 | -100.9838
19.7028 | -100.937
19.7065 | -100.8932
19.7103 | -100.8515
19.714 | -100.8124
19.7178 | -100.7781
19.7215 | -100.7476
19.7253 | -100.7208
19.729 | -100.695
19.7328 | -100.6712
19.7365 | -100.6496
19.7403 | -100.6337
19.744 | -100.622
19.7478 | -100.6124
19.7515 | -100.6055
19.7553 | -100.6008
19.759 | -100.5994
19.7628 | -100.6013
19.7665 | -100.6066
19.7703 | -100.6156
19.774 | -100.6298
19.7778 | -100.6472
19.7815 | -100.6664
19.7853 | -100.689
19.789 | -100.7174
19.7928 | -100.748
19.7965 | -100.7816
19.8003 | -100.8184
19.804 | -100.8585
19.8078 | -100.9021
19.8115 | -100.9491
19.8153 | -101.0004
19.819 | -101.0546
19.8228 | -101.1119
19.8265 | -101.1691
19.8303 | -101.227
19.834 | -101.2874
19.8378 | -101.3489
19.8415 | -101.4145
19.8453 | -101.4883
19.849 | -101.5622
19.8528 | -101.6301
19.8565 | -101.6932
19.8603 | -101.7541
19.864 | -101.8199
19.8678 | -101.8905
19.8715 | -101.9596
19.8753 | -102.0226
19.879 | -102.0779
19.8828 | -102.1287
19.8865 | -102.1785
19.8903 | -102.2277
19.894 | -102.2751
19.8978 | -102.3205
19.9015 | -102.3607
19.9053 | -102.3965
19.909 | -102.4294
19.9128 | -102.4636
19.9165 | -102.4989
19.9203 | -102.5327
19.924 | -102.561
19.9278 | -102.5793
19.9315 | -102.5931
19.9353 | -102.6099
19.9391 | -102.6328
19.9428 | -102.6603
19.9466 | -102.6781
19.9503 | -102.6821
19.9541 | -102.6736
19.9578 | -102.6625
19.9616 | -102.6576
19.9653 | -102.6542
19.9691 | -102.6471
19.9728 | -102.635
19.9766 | -102.6187
19.9803 | -102.5994
19.9841 | -102.5811
19.9878 | -102.562
19.9916 | -102.542
19.9953 | -102.5211
19.9991 | -102.4966
20.0028 | -102.4679
20.0066 | -102.4373
20.0103 | -102.4076
20.0141 | -102.3788
20.0178 | -102.3345
20.0216 | -102.3034
20.0253 | -102.2691
20.0291 | -102.2327
20.0328 | -102.1947
20.0366 | -102.1556
20.0403 | -102.1179
20.0441 | -102.0793
20.0478 | -102.0394
20.0516 | -101.9998
20.0553 | -101.9612
20.0591 | -101.9219
20.0628 | -101.8813
20.0666 | -101.8414
20.0703 | -101.8023
20.0741 | -101.7643
20.0778 | -101.728
20.0816 | -101.696
20.0853 | -101.6667
20.0891 | -101.6374
20.0928 | -101.6106
20.0966 | -101.5868
20.1003 | -101.5653
20.1041 | -101.5463
20.1078 | -101.5303
20.1116 | -101.5177
20.1153 | -101.5076
20.1191 | -101.4976
20.1228 | -101.4916
20.1266 | -101.4887
20.1303 | -101.4888
20.1341 | -101.4933
20.1378 | -101.5
20.1416 | -101.5098
20.1453 | -101.5234
20.1491 | -101.5389
20.1528 | -101.5562
20.1566 | -101.5775
20.1603 | -101.601
20.1641 | -101.6268
20.1678 | -101.6555
20.1716 | -101.6859
20.1753 | -101.7162
20.1791 | -101.7459
20.1828 | -101.7786
20.1866 | -101.8135
20.1904 | -101.8481
20.1941 | -101.884
20.1979 | -101.9192
20.2016 | -101.952
20.2054 | -101.9835
20.2091 | -102.016
20.2129 | -102.048
20.2166 | -102.0769
20.2204 | -102.1051
20.2241 | -102.1305
20.2279 | -102.151
20.2316 | -102.1698
20.2354 | -102.1845
20.2391 | -102.1962
20.2429 | -102.2037
20.2466 | -102.2053
20.2504 | -102.2028
20.2541 | -102.1986
20.2579 | -102.1924
20.2616 | -102.1812
20.2654 | -102.1637
20.2691 | -102.1432
20.2729 | -102.1175
20.2766 | -102.0868
20.2804 | -102.0549
20.2841 | -102.0231
20.2879 | -101.9892
20.2916 | -101.9528
20.2954 | -101.9148
20.2991 | -101.8738
20.3029 | -101.8325
20.3066 | -101.7924
20.3104 | -101.7507
20.3141 | -101.7104
20.3179 | -101.6705
20.3216 | -101.6316
20.3254 | -101.5944
20.3291 | -101.5589
20.3329 | -101.5269
20.3366 | -101.4979
20.3404 | -101.4705
20.3441 | -101.446
20.3479 | -101.4252
20.3516 | -101.4092
20.3554 | -101.3978
20.3591 | -101.3898
20.3629 | -101.3849
20.3666 | -101.383
20.3704 | -101.3837
20.3741 | -101.388
20.3779 | -101.3979
20.3816 | -101.4113
20.3854 | -101.4293
20.3891 | -101.452
20.3929 | -101.4786
20.3966 | -101.5092
20.4004 | -101.545
20.4041 | -101.5847
20.4079 | -101.6285
20.4116 | -101.675
20.4154 | -101.7228
20.4191 | -101.7709
20.4229 | -101.8217
20.4266 | -101.8756
20.4304 | -101.934
20.4342 | -101.9932
20.4379 | -102.0515
20.4417 | -102.1113
20.4454 | -102.1705
20.4492 | -102.2289
20.4529 | -102.2859
20.4567 | -102.3408
20.4604 | -102.3941
20.4642 | -102.4446
20.4679 | -102.4926
20.4717 | -102.5367
20.4754 | -102.5748
20.4792 | -102.61
20.4829 | -102.6395
20.4867 | -102.6616
20.4904 | -102.6799
20.4942 | -102.6955
20.4979 | -102.7034
20.5017 | -102.7071
20.5054 | -102.7054
20.5092 | -102.6973
20.5129 | -102.6871
20.5167 | -102.6708
20.5204 | -102.6493
20.5242 | -102.6239
20.5279 | -102.5952
20.5317 | -102.5629
20.5354 | -102.5268
20.5392 | -102.4883
20.5429 | -102.4456
20.5467 | -102.4011
20.5504 | -102.3581
20.5542 | -102.3191
20.5579 | -102.2779
20.5617 | -102.2424
20.5654 | -102.2048
20.5692 | -102.1687
20.5729 | -102.13
20.5767 | -102.0908
20.5804 | -102.0531
20.5842 | -102.0179
20.5879 | -101.9873
20.5917 | -101.9625
20.5954 | -101.9409
20.5992 | -101.9204
20.6029 | -101.9023
20.6067 | -101.8862
20.6104 | -101.8735
20.6142 | -101.8649
20.6179 | -101.8582
20.6217 | -101.8525
20.6254 | -101.8492
20.6292 | -101.8508
20.6329 | -101.855
20.6367 | -101.8578
20.6404 | -101.8639
20.6442 | -101.8735
20.6479 | -101.8839
20.6517 | -101.8942
20.6554 | -101.9047
20.6592 | -101.9178
20.6629 | -101.9325
20.6667 | -101.9479
20.6704 | -101.9616
20.6742 | -101.9737
20.6779 | -101.985
20.6817 | -101.9964
20.6855 | -102.0068
20.6892 | -102.0148
20.693 | -102.0195
20.6967 | -102.0196
20.7005 | -102.0191
20.7042 | -102.0201
20.708 | -102.0204
20.7117 | -102.0182
20.7155 | -102.013
20.7192 | -102.0039
20.723 | -101.9926
20.7267 | -101.9812
20.7305 | -101.9679
20.7342 | -101.9536
20.738 | -101.9384
20.7417 | -101.9224
20.7455 | -101.9062
20.7492 | -101.8883
20.753 | -101.8696
20.7567 | -101.8524
20.7605 | -101.8335
20.7642 | -101.8109
20.768 | -101.7879
20.7717 | -101.7639
20.7755 | -101.7403
20.7792 | -101.719
20.783 | -101.6984
20.7867 | -101.6781
20.7905 | -101.6581
20.7942 | -101.6355
20.798 | -101.6158
20.8017 | -101.5995
20.8055 | -101.583
20.8092 | -101.5683
20.813 | -101.556
20.8167 | -101.548
20.8205 | -101.5424
20.8242 | -101.5382
20.828 | -101.5337
20.8317 | -101.5309
20.8355 | -101.5333
20.8392 | -101.5384
20.843 | -101.5447
20.8467 | -101.5526
20.8505 | -101.5619
20.8542 | -101.574
20.858 | -101.5901
20.8617 | -101.6085
20.8655 | -101.6298
20.8692 | -101.6555
20.873 | -101.6837
20.8767 | -101.7145
20.8805 | -101.7474
20.8842 | -101.7804
20.888 | -101.8136
20.8917 | -101.8497
20.8955 | -101.8893
20.8992 | -101.9324
20.903 | -101.9788
20.9067 | -102.0256
20.9105 | -102.07
20.9142 | -102.1164
20.918 | -102.1661
20.9217 | -102.2177
20.9255 | -102.2666
20.9293 | -102.3173
20.933 | -102.3707
20.9368 | -102.4206
20.9405 | -102.4706
20.9443 | -102.5218
20.948 | -102.5748
20.9518 | -102.6273
20.9555 | -102.6774
20.9593 | -102.7249
20.963 | -102.771
20.9668 | -102.8153
20.9705 | -102.8598
20.9743 | -102.9025
20.978 | -102.9408
20.9818 | -102.9775
20.9855 | -103.0117
20.9893 | -103.0409
20.993 | -103.0713
20.9968 | -103.0989
21.0005 | -103.12
21.0043 | -103.1387
21.008 | -103.1557
21.0118 | -103.1673
21.0155 | -103.1749
21.0193 | -103.176
21.023 | -103.1762
21.0268 | -103.1767
21.0305 | -103.1732
21.0343 | -103.1641
21.038 | -103.1516
21.0418 | -103.138
21.0455 | -103.1246
21.0493 | -103.1095
21.053 | -103.0913
21.0568 | -103.0693
21.0605 | -103.0482
21.0643 | -103.0278
21.068 | -103.0054
21.0718 | -102.982
21.0755 | -102.955
21.0793 | -102.9277
21.083 | -102.8992
21.0868 | -102.8726
21.0905 | -102.8503
21.0943 | -102.8262
21.098 | -102.8004
21.1018 | -102.7829
21.1055 | -102.7604
21.1093 | -102.7405
21.113 | -102.7206
21.1168 | -102.7003
21.1205 | -102.6826
21.1243 | -102.6678
21.128 | -102.6562
21.1318 | -102.646
21.1355 | -102.636
21.1393 | -102.6271
21.143 | -102.6206
21.1468 | -102.6169
21.1505 | -102.616
21.1543 | -102.6151
21.158 | -102.6167
21.1618 | -102.6213
21.1655 | -102.626
21.1693 | -102.633
21.1731 | -102.6422
21.1768 | -102.6539
21.1806 | -102.6662
21.1843 | -102.6803
21.1881 | -102.6958
21.1918 | -102.7115
21.1956 | -102.7283
21.1993 | -102.7451
21.2031 | -102.7631
21.2068 | -102.7824
21.2106 | -102.8061
21.2143 | -102.83
21.2181 | -102.8533
21.2218 | -102.8756
21.2256 | -102.8952
21.2293 | -102.9137
21.2331 | -102.9324
21.2368 | -102.9532
21.2406 | -102.9723
21.2443 | -102.9901
21.2481 | -103.0039
21.2518 | -103.0109
21.2556 | -103.0157
21.2593 | -103.0212
21.2631 | -103.0261
21.2668 | -103.0335
21.2706 | -103.0364
21.2743 | -103.0318
21.2781 | -103.0247
21.2818 | -103.0125
21.2856 | -102.999
21.2893 | -102.9895
21.2931 | -102.977
21.2968 | -102.9592
21.3006 | -102.9378
21.3043 | -102.9128
21.3081 | -102.886
21.3118 | -102.8611
21.3156 | -102.8363
21.3193 | -102.811
21.3231 | -102.7793
21.3268 | -102.7382
21.3306 | -102.6962
21.3343 | -102.6578
21.3381 | -102.6235
21.3418 | -102.5874
21.3456 | -102.5496
21.3493 | -102.511
21.3531 | -102.4724
21.3568 | -102.4368
21.3606 | -102.4025
21.3643 | -102.3678
21.3681 | -102.3343
21.3718 | -102.3015
21.3756 | -102.2709
21.3793 | -102.2397
21.3831 | -102.2093
21.3868 | -102.1809
21.3906 | -102.1529
21.3943 | -102.1274
21.3981 | -102.1035
21.4018 | -102.0803
21.4056 | -102.0609
21.4093 | -102.0437
21.4131 | -102.0285
21.4169 | -102.0171
21.4206 | -102.0076
21.4244 | -102.0014
21.4281 | -101.9967
21.4319 | -101.9944
21.4356 | -101.9953
21.4394 | -102.0003
21.4431 | -102.0094
21.4469 | -102.0207
21.4506 | -102.0341
21.4544 | -102.0505
21.4581 | -102.0703
21.4619 | -102.0922
21.4656 | -102.1157
21.4694 | -102.1438
21.4731 | -102.1728
21.4769 | -102.2025
21.4806 | -102.2345
21.4844 | -102.2705
21.4881 | -102.3075
21.4919 | -102.3436
21.4956 | -102.3799
21.4994 | -102.4159
21.5031 | -102.4526
21.5069 | -102.4927
21.5106 | -102.535
21.5144 | -102.5766
21.5181 | -102.6161
21.5219 | -102.6538
21.5256 | -102.6914
21.5294 | -102.7327
21.5331 | -102.7734
21.5369 | -102.8108
21.5406 | -102.8436
21.5444 | -102.877
21.5481 | -102.9101
21.5519 | -102.9431
21.5556 | -102.9789
21.5594 | -103.012
21.5631 | -103.039
21.5669 | -103.0624
21.5706 | -103.0814
21.5744 | -103.0991
21.5781 | -103.119
21.5819 | -103.1369
21.5856 | -103.1536
21.5894 | -103.1656
21.5931 | -103.1693
21.5969 | -103.1676
21.6006 | -103.167
21.6044 | -103.1676
21.6081 | -103.168
21.6119 | -103.1666
21.6156 | -103.1603
21.6194 | -103.1518
21.6231 | -103.1433
21.6269 | -103.1387
21.6306 | -103.1393
21.6344 | -103.1386
21.6381 | -103.133
21.6419 | -103.1227
21.6456 | -103.1107
21.6494 | -103.1
21.6531 | -103.091
21.6569 | -103.0818
21.6607 | -103.0697
21.6644 | -103.0579
21.6682 | -103.046
21.6719 | -103.0352
21.6757 | -103.0267
21.6794 | -103.0195
21.6832 | -103.0118
21.6869 | -103.0034
21.6907 | -102.9964
21.6944 | -102.9902
21.6982 | -102.9859
21.7019 | -102.9828
21.7057 | -102.981
21.7094 | -102.9789
21.7132 | -102.9747
21.7169 | -102.9709
21.7207 | -102.9658
21.7244 | -102.9599
21.7282 | -102.9555
21.7319 | -102.9484
21.7357 | -102.9416
21.7394 | -102.9369
21.7432 | -102.9332
21.7469 | -102.9256
21.7507 | -102.9161
21.7544 | -102.9101
21.7582 | -102.9048
21.7619 | -102.8984
21.7657 | -102.8924
21.7694 | -102.8854
21.7732 | -102.8782
21.7769 | -102.8716
21.7807 | -102.8665
21.7844 | -102.8621
21.7882 | -102.8574
21.7919 | -102.8507
21.7957 | -102.846
21.7994 | -102.8451
21.8032 | -102.8451
21.8069 | -102.8454
21.8107 | -102.844
21.8144 | -102.8374
21.8182 | -102.8312
21.8219 | -102.8287
21.8257 | -102.8286
21.8294 | -102.8273
21.8332 | -102.8262
21.8369 | -102.8257
21.8407 | -102.825
21.8444 | -102.8261
21.8482 | -102.8272
21.8519 | -102.8276
21.8557 | -102.8271
21.8594 | -102.8255
21.8632 | -102.8259
21.8669 | -102.8261
21.8707 | -102.8282
21.8744 | -102.8292
21.8782 | -102.8285
21.8819 | -102.83
21.8857 | -102.8303
21.8894 | -102.8291
21.8932 | -102.8277
21.8969 | -102.8273
21.9007 | -102.8277
21.9044 | -102.8302
21.9082 | -102.8305
21.912 | -102.8276
21.9157 | -102.8253
21.9195 | -102.8225
21.9232 | -102.8189
21.927 | -102.8161
21.9307 | -102.8145
21.9345 | -102.811
21.9382 | -102.8071
21.942 | -102.8021
21.9457 | -102.7992
21.9495 | -102.7979
21.9532 | -102.7988
21.957 | -102.7986
21.9607 | -102.7977
21.9645 | -102.7976
21.9682 | -102.7981
21.972 | -102.8014
21.9757 | -102.8053
21.9795 | -102.8097
21.9832 | -102.812
21.987 | -102.8123
21.9907 | -102.8136
21.9945 | -102.8172
21.9982 | -102.8231
22.002 | -102.8271
22.0057 | -102.8313
22.0095 | -102.8353
22.0132 | -102.8388
22.017 | -102.84
22.0207 | -102.8402
22.0245 | -102.843
22.0282 | -102.8458
22.032 | -102.8513
22.0357 | -102.8578
22.0395 | -102.8632
22.0432 | -102.8674
22.047 | -102.8716
22.0507 | -102.8744
22.0545 | -102.8735
22.0582 | -102.8735
22.062 | -102.8771
22.0657 | -102.8768
22.0695 | -102.8723
22.0732 | -102.8652
22.077 | -102.8619
22.0807 | -102.8585
22.0845 | -102.8553
22.0882 | -102.8509
22.092 | -102.8464
22.0957 | -102.8409
22.0995 | -102.8321
22.1032 | -102.8242
22.107 | -102.8148
22.1107 | -102.8032
22.1145 | -102.7923
22.1182 | -102.7812
22.122 | -102.7708
22.1257 | -102.7579
22.1295 | -102.7405
22.1332 | -102.7217
22.137 | -102.7051
22.1407 | -102.6901
22.1445 | -102.6693
22.1482 | -102.6459
22.152 | -102.6244
22.1558 | -102.6038
22.1595 | -102.5862
22.1633 | -102.568
22.167 | -102.5495
22.1708 | -102.532
22.1745 | -102.5151
22.1783 | -102.4997
22.182 | -102.4844
22.1858 | -102.4746
22.1895 | -102.4608
22.1933 | -102.4491
22.197 | -102.4365
22.2008 | -102.4257
22.2045 | -102.4169
22.2083 | -102.4089
22.212 | -102.4022
22.2158 | -102.3953
22.2195 | -102.3875
22.2233 | -102.3817
22.227 | -102.3802
22.2308 | -102.3804
22.2345 | -102.3817
22.2383 | -102.3835
22.242 | -102.3853
22.2458 | -102.3871
22.2495 | -102.3901
22.2533 | -102.3948
22.257 | -102.4011
22.2608 | -102.4074
22.2645 | -102.4151
22.2683 | -102.4242
22.272 | -102.4351
22.2758 | -102.4476
22.2795 | -102.4593
22.2833 | -102.4706
22.287 | -102.482
22.2908 | -102.4939
22.2945 | -102.5067
22.2983 | -102.5209
22.302 | -102.5322
22.3058 | -102.5433
22.3095 | -102.5548
22.3133 | -102.5673
22.317 | -102.5804
22.3208 | -102.5918
22.3245 | -102.6049
22.3283 | -102.6186
22.332 | -102.6343
22.3358 | -102.651
22.3395 | -102.665
22.3433 | -102.6785
22.347 | -102.6902
22.3508 | -102.7007
22.3545 | -102.7127
22.3583 | -102.7253
22.362 | -102.7379
22.3658 | -102.7492
22.3695 | -102.7615
22.3733 | -102.7744
22.377 | -102.7839
22.3808 | -102.7917
22.3845 | -102.8014
22.3883 | -102.8118
22.392 | -102.817
22.3958 | -102.8236
22.3996 | -102.8301
22.4033 | -102.8349
22.4071 | -102.838
22.4108 | -102.8411
22.4146 | -102.8453
22.4183 | -102.8462
22.4221 | -102.8473
22.4258 | -102.8503
22.4296 | -102.8518
22.4333 | -102.8507
22.4371 | -102.8468
22.4408 | -102.845
22.4446 | -102.8467
22.4483 | -102.8476
22.4521 | -102.8444
22.4558 | -102.8393
22.4596 | -102.8383
22.4633 | -102.8358
22.4671 | -102.8249
22.4708 | -102.8101
22.4746 | -102.7955
22.4783 | -102.7787
22.4821 | -102.76
22.4858 | -102.74
22.4896 | -102.716
22.4933 | -102.6921
22.4971 | -102.6636
22.5008 | -102.6306
22.5046 | -102.5945
22.5083 | -102.562
22.5121 | -102.5303
22.5158 | -102.497
22.5196 | -102.4605
22.5233 | -102.4232
22.5271 | -102.3841
22.5308 | -102.3418
22.5346 | -102.301
22.5383 | -102.2594
22.5421 | -102.2142
22.5458 | -102.1645
22.5496 | -102.1133
22.5533 | -102.0634
22.5571 | -102.0112
22.5608 | -101.9573
22.5646 | -101.9055
22.5683 | -101.8548
22.5721 | -101.8042
22.5758 | -101.754
22.5796 | -101.7033
22.5833 | -101.6537
22.5871 | -101.6034
22.5908 | -101.5547
22.5946 | -101.5089
22.5983 | -101.4665
22.6021 | -101.4279
22.6058 | -101.3915
22.6096 | -101.3564
22.6133 | -101.3222
22.6171 | -101.2912
22.6208 | -101.266
22.6246 | -101.2444
22.6283 | -101.2228
22.6321 | -101.2038
22.6358 | -101.1902
22.6396 | -101.181
22.6433 | -101.1737
22.6471 | -101.1696
22.6509 | -101.1684
22.6546 | -101.1695
22.6584 | -101.1734
22.6621 | -101.1789
22.6659 | -101.1876
22.6696 | -101.2009
22.6734 | -101.2156
22.6771 | -101.2313
22.6809 | -101.2499
22.6846 | -101.2719
22.6884 | -101.2957
22.6921 | -101.3208
22.6959 | -101.3462
22.6996 | -101.3741
22.7034 | -101.4032
22.7071 | -101.4326
22.7109 | -101.4623
22.7146 | -101.4926
22.7184 | -101.5244
22.7221 | -101.555
22.7259 | -101.5804
22.7296 | -101.6097
22.7334 | -101.6397
22.7371 | -101.6692
22.7409 | -101.699
22.7446 | -101.7253
22.7484 | -101.7487
22.7521 | -101.7712
22.7559 | -101.795
22.7596 | -101.8188
22.7634 | -101.8399
22.7671 | -101.8588
22.7709 | -101.8739
22.7746 | -101.8864
22.7784 | -101.9002
22.7821 | -101.9144
22.7859 | -101.927
22.7896 | -101.942
22.7934 | -101.9564
22.7971 | -101.9652
22.8009 | -101.9716
22.8046 | -101.9774
22.8084 | -101.9833
22.8121 | -101.9888
22.8159 | -101.9896
22.8196 | -101.9922
22.8234 | -101.9989
22.8271 | -102.0043
22.8309 | -102.002
22.8346 | -101.998
22.8384 | -101.9942
22.8421 | -101.9896
22.8459 | -101.9831
22.8496 | -101.9779
22.8534 | -101.9753
22.8571 | -101.9721
22.8609 | -101.9696
22.8646 | -101.9677
22.8684 | -101.9661
22.8721 | -101.9658
22.8759 | -101.9647
22.8796 | -101.968
22.8834 | -101.9744
22.8871 | -101.9769
22.8909 | -101.9785
22.8946 | -101.9813
22.8984 | -101.9847
22.9022 | -101.9895
22.9059 | -101.9937
22.9097 | -101.9995
22.9134 | -102.0051
22.9172 | -102.0101
22.9209 | -102.0146
22.9247 | -102.0196
22.9284 | -102.0278
22.9322 | -102.0369
22.9359 | -102.0475
22.9397 | -102.0566
22.9434 | -102.0648
22.9472 | -102.0721
22.9509 | -102.0826
22.9547 | -102.0926
22.9584 | -102.0998
22.9622 | -102.1068
22.9659 | -102.1148
22.9697 | -102.1227
22.9734 | -102.1296
22.9772 | -102.1386
22.9809 | -102.1462
22.9847 | -102.1542
22.9884 | -102.1604
22.9922 | -102.1635
22.9959 | -102.1675
22.9997 | -102.1748
23.0034 | -102.1803
23.0072 | -102.1854
23.0109 | -102.1917
23.0147 | -102.196
23.0184 | -102.1993
23.0222 | -102.202
23.0259 | -102.2052
23.0297 | -102.2101
23.0334 | -102.2127
23.0372 | -102.2097
23.0409 | -102.2018
23.0447 | -102.1958
23.0484 | -102.1911
23.0522 | -102.1861
23.0559 | -102.1816
23.0597 | -102.1766
23.0634 | -102.1684
23.0672 | -102.1575
23.0709 | -102.1518
23.0747 | -102.146
23.0784 | -102.1374
23.0822 | -102.1267
23.0859 | -102.1139
23.0897 | -102.0978
23.0934 | -102.0792
23.0972 | -102.0633
23.1009 | -102.0468
23.1047 | -102.0284
23.1084 | -102.0055
23.1122 | -101.9816
23.1159 | -101.9577
23.1197 | -101.9328
23.1234 | -101.9096
23.1272 | -101.8861
23.1309 | -101.8627
23.1347 | -101.8366
23.1384 | -101.8078
23.1422 | -101.7811
23.146 | -101.7545
23.1497 | -101.7266
23.1535 | -101.6989
23.1572 | -101.6727
23.161 | -101.6457
23.1647 | -101.6186
23.1685 | -101.594
23.1722 | -101.5701
23.176 | -101.5459
23.1797 | -101.5223
23.1835 | -101.5
23.1872 | -101.4791
23.191 | -101.4608
23.1947 | -101.4445
23.1985 | -101.4282
23.2022 | -101.4147
23.206 | -101.4018
23.2097 | -101.3874
23.2135 | -101.3735
23.2172 | -101.3595
23.221 | -101.3465
23.2247 | -101.3368
23.2285 | -101.3277
23.2322 | -101.3203
23.236 | -101.3145
23.2397 | -101.3095
23.2435 | -101.305
23.2472 | -101.3044
23.251 | -101.3038
23.2547 | -101.3023
23.2585 | -101.3017
23.2622 | -101.3024
23.266 | -101.3063
23.2697 | -101.3074
23.2735 | -101.3087
23.2772 | -101.3076
23.281 | -101.3062
23.2847 | -101.3049
23.2885 | -101.3042
23.2922 | -101.3012
23.296 | -101.297
23.2997 | -101.2931
23.3035 | -101.2886
23.3072 | -101.2818
23.311 | -101.2769
23.3147 | -101.2721
23.3185 | -101.2668
23.3222 | -101.2609
23.326 | -101.2533
23.3297 | -101.2457
23.3335 | -101.2361
23.3372 | -101.2243
23.341 | -101.2128
23.3447 | -101.2023
23.3485 | -101.1915
23.3522 | -101.1799
23.356 | -101.1663
23.3597 | -101.1509
23.3635 | -101.1381
23.3672 | -101.1275
23.371 | -101.1189
23.3747 | -101.1081
23.3785 | -101.0941
23.3822 | -101.0802
23.386 | -101.0655
23.3898 | -101.0504
23.3935 | -101.037
23.3973 | -101.0241
23.401 | -101.0094
23.4048 | -100.9954
23.4085 | -100.9824
23.4123 | -100.9718
23.416 | -100.9611
23.4198 | -100.9526
23.4235 | -100.9442
23.4273 | -100.9334
23.431 | -100.9258
23.4348 | -100.918
23.4385 | -100.9113
23.4423 | -100.9045
23.446 | -100.8955
23.4498 | -100.8879
23.4535 | -100.8793
23.4573 | -100.8727
23.461 | -100.8677
23.4648 | -100.8651
23.4685 | -100.862
23.4723 | -100.8587
23.476 | -100.8551
23.4798 | -100.8527
23.4835 | -100.8526
23.4873 | -100.8527
23.491 | -100.8538
23.4948 | -100.8563
23.4985 | -100.8579
23.5023 | -100.8589
23.506 | -100.8595
23.5098 | -100.8604
23.5135 | -100.862
23.5173 | -100.8651
23.521 | -100.8678
23.5248 | -100.8705
23.5285 | -100.8731
23.5323 | -100.875
23.536 | -100.877
23.5398 | -100.8808
23.5435 | -100.8858
23.5473 | -100.8899
23.551 | -100.8932
23.5548 | -100.8965
23.5585 | -100.9015
23.5623 | -100.9066
23.566 | -100.9109
23.5698 | -100.9142
23.5735 | -100.9191
23.5773 | -100.9209
23.581 | -100.923
23.5848 | -100.9279
23.5885 | -100.9335
23.5923 | -100.9392
23.596 | -100.947
23.5998 | -100.954
23.6035 | -100.9617
23.6073 | -100.9716
23.611 | -100.9836
23.6148 | -100.9961
23.6185 | -101.007
23.6223 | -101.0167
23.626 | -101.0277
23.6298 | -101.0411
23.6336 | -101.0531
23.6373 | -101.0655
23.6411 | -101.0783
23.6448 | -101.0885
23.6486 | -101.0975
23.6523 | -101.1055
23.6561 | -101.1145
23.6598 | -101.1254
23.6636 | -101.1329
23.6673 | -101.1406
23.6711 | -101.1505
23.6748 | -101.1583
23.6786 | -101.1638
23.6823 | -101.1711
23.6861 | -101.1766
23.6898 | -101.1846
23.6936 | -101.1927
23.6973 | -101.195
23.7011 | -101.1974
23.7048 | -101.2018
23.7086 | -101.2063
23.7123 | -101.2121
23.7161 | -101.2146
23.7198 | -101.216
23.7236 | -101.2177
23.7273 | -101.2201
23.7311 | -101.2241
23.7348 | -101.2257
23.7386 | -101.227
23.7423 | -101.2262
23.7461 | -101.2231
23.7498 | -101.2235
23.7536 | -101.224
23.7573 | -101.2236
23.7611 | -101.2221
23.7648 | -101.2197
23.7686 | -101.218
23.7723 | -101.2168
23.7761 | -101.2168
23.7798 | -101.2186
23.7836 | -101.2182
23.7873 | -101.218
23.7911 | -101.2166
23.7948 | -101.2136
23.7986 | -101.208
23.8023 | -101.2057
23.8061 | -101.2052
23.8098 | -101.2023
23.8136 | -101.1993
23.8173 | -101.1975
23.8211 | -101.1948
23.8248 | -101.1916
23.8286 | -101.1872
23.8323 | -101.1833
23.8361 | -101.1815
23.8398 | -101.1809
23.8436 | -101.1822
23.8473 | -101.1835
23.8511 | -101.1845
23.8548 | -101.1842
23.8586 | -101.1824
23.8623 | -101.1834
23.8661 | -101.1876
23.8698 | -101.1929
23.8736 | -101.1994
23.8773 | -101.204
23.8811 | -101.2082
23.8849 | -101.2145
23.8886 | -101.2208
23.8924 | -101.2273
23.8961 | -101.2335
23.8999 | -101.2398
23.9036 | -101.245
23.9074 | -101.2518
23.9111 | -101.2595
23.9149 | -101.269
23.9186 | -101.2802
23.9224 | -101.2898
23.9261 | -101.2993
23.9299 | -101.3084
23.9336 | -101.3165
23.9374 | -101.3265
23.9411 | -101.3367
23.9449 | -101.3446
23.9486 | -101.353
23.9524 | -101.3603
23.9561 | -101.367
23.9599 | -101.3754
23.9636 | -101.3852
23.9674 | -101.3912
23.9711 | -101.3973
23.9749 | -101.4043
23.9786 | -101.4107
23.9824 | -101.4186
23.9861 | -101.4253
23.9899 | -101.4321
23.9936 | -101.437
23.9974 | -101.4395
24.0011 | -101.4438
24.0049 | -101.4481
24.0086 | -101.4497
24.0124 | -101.4536
24.0161 | -101.4588
24.0199 | -101.4622
24.0236 | -101.4656
24.0274 | -101.4652
24.0311 | -101.4631
24.0349 | -101.4657
24.0386 | -101.4667
24.0424 | -101.4647
24.0461 | -101.4621
24.0499 | -101.4572
24.0536 | -101.4526
24.0574 | -101.4488
24.0611 | -101.4452
24.0649 | -101.4425
24.0686 | -101.4383
24.0724 | -101.4349
24.0761 | -101.4316
24.0799 | -101.4305
24.0836 | -101.4286
24.0874 | -101.4243
24.0911 | -101.4188
24.0949 | -101.4142
24.0986 | -101.4134
24.1024 | -101.4126
24.1061 | -101.4095
24.1099 | -101.406
24.1136 | -101.4013
24.1174 | -101.3951
24.1212 | -101.3903
24.1249 | -101.3841
24.1287 | -101.3794
24.1324 | -101.3745
24.1362 | -101.372
24.1399 | -101.3681
24.1437 | -101.363
24.1474 | -101.3575
24.1512 | -101.3536
24.1549 | -101.3492
24.1587 | -101.3434
24.1624 | -101.3391
24.1662 | -101.3354
24.1699 | -101.3293
24.1737 | -101.3225
24.1774 | -101.318
24.1812 | -101.3145
24.1849 | -101.3104
24.1887 | -101.3066
24.1924 | -101.3039
24.1962 | -101.3019
24.1999 | -101.2996
24.2037 | -101.2957
24.2074 | -101.2917
24.2112 | -101.2865
24.2149 | -101.2833
24.2187 | -101.2828
24.2224 | -101.2822
24.2262 | -101.2818
24.2299 | -101.2814
24.2337 | -101.2792
24.2374 | -101.278
24.2412 | -101.2763
24.2449 | -101.2757
24.2487 | -101.2732
24.2524 | -101.2696
24.2562 | -101.2653
24.2599 | -101.2531
24.2637 | -101.2349
24.2674 | -101.2211
24.2712 | -101.2135
24.2749 | -101.2086
24.2787 | -101.2034
24.2824 | -101.1964
24.2862 | -101.1918
24.2899 | -101.1876
24.2937 | -101.1829
24.2974 | -101.1796
24.3012 | -101.1738
24.3049 | -101.1671
24.3087 | -101.1609
24.3124 | -101.1549
24.3162 | -101.1475
24.3199 | -101.1375
24.3237 | -101.1263
24.3274 | -101.1171
24.3312 | -101.1093
24.3349 | -101.1019
24.3387 | -101.0953
24.3424 | -101.088
24.3462 | -101.0796
24.3499 | -101.0713
24.3537 | -101.0615
24.3574 | -101.0496
24.3612 | -101.0384
24.3649 | -101.0286
24.3687 | -101.0212
24.3725 | -101.0144
24.3762 | -101.0095
24.38 | -101.0028
24.3837 | -100.9971
24.3875 | -100.9893
24.3912 | -100.9833
24.395 | -100.9814
24.3987 | -100.9814
24.4025 | -100.981
24.4062 | -100.9802
24.41 | -100.9788
24.4137 | -100.9786
24.4175 | -100.9796
24.4212 | -100.982
24.425 | -100.9851
24.4287 | -100.9891
24.4325 | -100.9936
24.4362 | -100.9999
24.44 | -101.0064
24.4437 | -101.0128
24.4475 | -101.0189
24.4512 | -101.0274
24.455 | -101.039
24.4587 | -101.0526
24.4625 | -101.0678
24.4662 | -101.0826
24.47 | -101.0992
24.4737 | -101.1153
24.4775 | -101.1315
24.4812 | -101.1476
24.485 | -101.1651
24.4887 | -101.1812
24.4925 | -101.2004
24.4962 | -101.2181
24.5 | -101.236
24.5037 | -101.2542
24.5075 | -101.2724
24.5112 | -101.2918
24.515 | -101.314
24.5187 | -101.3368
24.5225 | -101.3558
24.5262 | -101.373
24.53 | -101.3908
24.5337 | -101.4093
24.5375 | -101.4265
24.5412 | -101.4447
24.545 | -101.4649
24.5487 | -101.4848
24.5525 | -101.5025
24.5562 | -101.523
24.56 | -101.5403
24.5637 | -101.5573
24.5675 | -101.5752
24.5712 | -101.5932
24.575 | -101.6096
24.5787 | -101.6214
24.5825 | -101.6203
24.5862 | -101.6105
24.59 | -101.5939
24.5937 | -101.587
24.5975 | -101.5862
24.6012 | -101.5867
24.605 | -101.5932
24.6087 | -101.603
24.6125 | -101.6018
24.6163 | -101.5909
24.62 | -101.5743
24.6238 | -101.5563
24.6275 | -101.5395
24.6313 | -101.5192
24.635 | -101.4965
24.6388 | -101.4718
24.6425 | -101.4474
24.6463 | -101.4226
24.65 | -101.3952
24.6538 | -101.3662
24.6575 | -101.3352
24.6613 | -101.3035
24.665 | -101.2735
24.6688 | -101.242
24.6725 | -101.2076
24.6763 | -101.1707
24.68 | -101.1319
24.6838 | -101.0926
24.6875 | -101.0536
24.6913 | -101.0174
24.695 | -100.9821
24.6988 | -100.9454
24.7025 | -100.9105
24.7063 | -100.8772
24.71 | -100.8444
24.7138 | -100.8125
24.7175 | -100.7812
24.7213 | -100.7506
24.725 | -100.7225
24.7288 | -100.6967
24.7325 | -100.6721
24.7363 | -100.6525
24.74 | -100.6349
24.7438 | -100.6178
24.7475 | -100.6037
24.7513 | -100.5913
24.755 | -100.5829
24.7588 | -100.5772
24.7625 | -100.5751
24.7663 | -100.5765
24.77 | -100.5811
24.7738 | -100.5899
24.7775 | -100.6032
24.7813 | -100.6195
24.785 | -100.6385
24.7888 | -100.6604
24.7925 | -100.6854
24.7963 | -100.7138
24.8 | -100.7455
24.8038 | -100.7779
24.8075 | -100.8104
24.8113 | -100.8446
24.815 | -100.8812
24.8188 | -100.9179
24.8225 | -100.957
24.8263 | -101
24.83 | -101.044
24.8338 | -101.0874
24.8375 | -101.1336
24.8413 | -101.1782
24.845 | -101.2222
24.8488 | -101.2682
24.8525 | -101.3129
24.8563 | -101.355
24.8601 | -101.3968
24.8638 | -101.4382
24.8676 | -101.4767
24.8713 | -101.5133
24.8751 | -101.549
24.8788 | -101.5807
24.8826 | -101.6103
24.8863 | -101.6412
24.8901 | -101.6484
24.8938 | -101.6726
24.8976 | -101.6939
24.9013 | -101.715
24.9051 | -101.735
24.9088 | -101.7509
24.9126 | -101.7656
24.9163 | -101.7756
24.9201 | -101.7856
24.9238 | -101.7923
24.9276 | -101.7972
24.9313 | -101.8032
24.9351 | -101.8193
24.9388 | -101.8398
24.9426 | -101.8619
24.9463 | -101.8684
24.9501 | -101.8628
24.9538 | -101.8546
24.9576 | -101.8455
24.9613 | -101.8355
24.9651 | -101.8269
24.9688 | -101.8191
24.9726 | -101.8135
24.9763 | -101.8065
24.9801 | -101.7992
24.9838 | -101.7926
24.9876 | -101.7856
24.9913 | -101.7792
24.9951 | -101.7715
24.9988 | -101.7637
25.0026 | -101.7578
25.0063 | -101.7543
25.0101 | -101.7523
25.0138 | -101.7502
25.0176 | -101.7482
25.0213 | -101.7466
25.0251 | -101.7443
25.0288 | -101.7443
25.0326 | -101.7476
25.0363 | -101.7516
25.0401 | -101.7554
25.0438 | -101.7581
25.0476 | -101.7632
25.0513 | -101.7724
25.0551 | -101.781
25.0588 | -101.7896
25.0626 | -101.7977
25.0663 | -101.8053
25.0701 | -101.8149
25.0738 | -101.8269
25.0776 | -101.8374
25.0813 | -101.8503
25.0851 | -101.8646
25.0888 | -101.8779
25.0926 | -101.8919
25.0963 | -101.906
25.1001 | -101.9211
25.1038 | -101.9372
25.1076 | -101.9519
25.1114 | -101.9688
25.1151 | -101.9846
25.1189 | -101.9997
25.1226 | -102.0152
25.1264 | -102.0296
25.1301 | -102.0432
25.1339 | -102.058
25.1376 | -102.0736
25.1414 | -102.0901
25.1451 | -102.1043
25.1489 | -102.1188
25.1526 | -102.1319
25.1564 | -102.1449
25.1601 | -102.1613
25.1639 | -102.1766
25.1676 | -102.1894
25.1714 | -102.2048
25.1751 | -102.2194
25.1789 | -102.232
25.1826 | -102.2441
25.1864 | -102.2562
25.1901 | -102.2653
25.1939 | -102.2771
25.1976 | -102.2876
25.2014 | -102.2946
25.2051 | -102.3039
25.2089 | -102.3126
25.2126 | -102.3191
25.2164 | -102.3255
25.2201 | -102.3338
25.2239 | -102.3431
25.2276 | -102.352
25.2314 | -102.3619
25.2351 | -102.3702
25.2389 | -102.3775
25.2426 | -102.3855
25.2464 | -102.3928
25.2501 | -102.4006
25.2539 | -102.4092
25.2576 | -102.4161
25.2614 | -102.423
25.2651 | -102.4295
25.2689 | -102.4348
25.2726 | -102.4427
25.2764 | -102.45
25.2801 | -102.4593
25.2839 | -102.4685
25.2876 | -102.477
25.2914 | -102.4841
25.2951 | -102.4904
25.2989 | -102.4959
25.3026 | -102.5025
25.3064 | -102.5098
25.3101 | -102.5151
25.3139 | -102.5205
25.3176 | -102.5264
25.3214 | -102.5353
25.3251 | -102.5421
25.3289 | -102.5482
25.3326 | -102.5549
25.3364 | -102.5606
25.3401 | -102.5653
25.3439 | -102.5716
25.3476 | -102.58
25.3514 | -102.5881
25.3551 | -102.5971
25.3589 | -102.6035
25.3627 | -102.6084
25.3664 | -102.6137
25.3702 | -102.6197
25.3739 | -102.6253
25.3777 | -102.6313
25.3814 | -102.6376
25.3852 | -102.6418
25.3889 | -102.645
25.3927 | -102.6488
25.3964 | -102.6551
25.4002 | -102.6624
25.4039 | -102.6668
25.4077 | -102.6699
25.4114 | -102.6751
25.4152 | -102.6795
25.4189 | -102.6853
25.4227 | -102.6915
25.4264 | -102.694
25.4302 | -102.6976
25.4339 | -102.7017
25.4377 | -102.7068
25.4414 | -102.7118
25.4452 | -102.7163
25.4489 | -102.722
25.4527 | -102.7282
25.4564 | -102.7341
25.4602 | -102.7401
25.4639 | -102.7471
25.4677 | -102.7543
25.4714 | -102.7606
25.4752 | -102.7685
25.4789 | -102.7773
25.4827 | -102.7835
25.4864 | -102.7919
25.4902 | -102.7979
25.4939 | -102.8036
25.4977 | -102.8089
25.5014 | -102.8149
25.5052 | -102.8228
25.5089 | -102.8285
25.5127 | -102.8351
25.5164 | -102.8424
25.5202 | -102.8478
25.5239 | -102.8538
25.5277 | -102.8614
25.5314 | -102.8672
25.5352 | -102.8682
25.5389 | -102.866
25.5427 | -102.8666
25.5464 | -102.8689
25.5502 | -102.8724
25.5539 | -102.8748
25.5577 | -102.8771
25.5614 | -102.8778
25.5652 | -102.8764
25.5689 | -102.8765
25.5727 | -102.8771
25.5764 | -102.8762
25.5802 | -102.8731
25.5839 | -102.871
25.5877 | -102.8683
25.5914 | -102.8642
25.5952 | -102.8589
25.5989 | -102.8512
25.6027 | -102.8451
25.6065 | -102.8406
25.6102 | -102.8363
25.614 | -102.8325
25.6177 | -102.8272
25.6215 | -102.8207
25.6252 | -102.8149
25.629 | -102.8121
25.6327 | -102.8091
25.6365 | -102.8033
25.6402 | -102.7961
25.644 | -102.7901
25.6477 | -102.7836
25.6515 | -102.7773
25.6552 | -102.7709
25.659 | -102.7649
25.6627 | -102.7601
25.6665 | -102.756
25.6702 | -102.7539
25.674 | -102.7508
25.6777 | -102.7488
25.6815 | -102.7458
25.6852 | -102.7458
25.689 | -102.7461
25.6927 | -102.7444
25.6965 | -102.7435
25.7002 | -102.7457
25.704 | -102.7485
25.7077 | -102.7512
25.7115 | -102.7541
25.7152 | -102.7555
25.719 | -102.7564
25.7227 | -102.7584
25.7265 | -102.7616
25.7302 | -102.7651
25.734 | -102.7686
25.7377 | -102.7731
25.7415 | -102.7794
25.7452 | -102.7842
25.749 | -102.7885
25.7527 | -102.7922
25.7565 | -102.7944
25.7602 | -102.7964
25.764 | -102.7982
25.7677 | -102.8013
25.7715 | -102.8036
25.7752 | -102.806
25.779 | -102.8088
25.7827 | -102.81
25.7865 | -102.812
25.7902 | -102.8117
25.794 | -102.8112
25.7977 | -102.8093
25.8015 | -102.8073
25.8052 | -102.8049
25.809 | -102.8014
25.8127 | -102.7961
25.8165 | -102.7916
25.8202 | -102.7858
25.824 | -102.7775
25.8277 | -102.7709
25.8315 | -102.7649
25.8352 | -102.758
25.839 | -102.7496
25.8427 | -102.7417
25.8465 | -102.7343
25.8503 | -102.7253
25.854 | -102.7187
25.8578 | -102.7135
25.8615 | -102.7058
25.8653 | -102.6969
25.869 | -102.6882
25.8728 | -102.6804
25.8765 | -102.6738
25.8803 | -102.6672
25.884 | -102.6616
25.8878 | -102.655
25.8915 | -102.6474
25.8953 | -102.6416
25.899 | -102.6385
25.9028 | -102.637
25.9065 | -102.6328
25.9103 | -102.6286
25.914 | -102.6244
25.9178 | -102.6203
25.9215 | -102.6197
25.9253 | -102.6204
25.929 | -102.6199
25.9328 | -102.6187
25.9365 | -102.6172
25.9403 | -102.6169
25.944 | -102.6173
25.9478 | -102.617
25.9515 | -102.6184
25.9553 | -102.6198
25.959 | -102.6215
25.9628 | -102.6238
25.9665 | -102.6266
25.9703 | -102.6307
25.974 | -102.6335
25.9778 | -102.6359
25.9815 | -102.6348
25.9853 | -102.6326
25.989 | -102.6344
25.9928 | -102.6384
25.9965 | -102.642
26.0003 | -102.6439
26.004 | -102.6478
26.0078 | -102.6518
26.0115 | -102.6551
26.0153 | -102.6586
26.019 | -102.6606
26.0228 | -102.6628
26.0265 | -102.664
26.0303 | -102.6639
26.034 | -102.6639
26.0378 | -102.663
26.0415 | -102.6648
26.0453 | -102.666
26.049 | -102.6665
26.0528 | -102.6667
26.0565 | -102.6678
26.0603 | -102.6701
26.064 | -102.6719
26.0678 | -102.6727
26.0715 | -102.6743
26.0753 | -102.6749
26.079 | -102.6762
26.0828 | -102.6793
26.0865 | -102.6842
26.0903 | -102.6881
26.094 | -102.6907
26.0978 | -102.6948
26.1016 | -102.6995
26.1053 | -102.7036
26.1091 | -102.7059
26.1128 | -102.71
26.1166 | -102.7162
26.1203 | -102.7224
26.1241 | -102.7286
26.1278 | -102.7355
26.1316 | -102.7447
26.1353 | -102.7525
26.1391 | -102.7607
26.1428 | -102.7713
26.1466 | -102.7822
26.1503 | -102.7946
26.1541 | -102.8076
26.1578 | -102.8218
26.1616 | -102.8354
26.1653 | -102.8507
26.1691 | -102.8666
26.1728 | -102.8813
26.1766 | -102.8962
26.1803 | -102.9104
26.1841 | -102.9245
26.1878 | -102.9433
26.1916 | -102.9622
26.1953 | -102.9787
26.1991 | -102.9955
26.2028 | -103.014
26.2066 | -103.0343
26.2103 | -103.0536
26.2141 | -103.074
26.2178 | -103.0946
26.2216 | -103.1149
26.2253 | -103.1331
26.2291 | -103.1516
26.2328 | -103.1705
26.2366 | -103.1882
26.2403 | -103.2051
26.2441 | -103.2211
26.2478 | -103.237
26.2516 | -103.2518
26.2553 | -103.2681
26.2591 | -103.2854
26.2628 | -103.2988
26.2666 | -103.3098
26.2703 | -103.3209
26.2741 | -103.331
26.2778 | -103.3373
26.2816 | -103.3424
26.2853 | -103.3487
26.2891 | -103.3536
26.2928 | -103.3566
26.2966 | -103.3607
26.3003 | -103.3647
26.3041 | -103.3697
26.3078 | -103.3734
26.3116 | -103.375
26.3153 | -103.3745
26.3191 | -103.3735
26.3228 | -103.3729
26.3266 | -103.3699
26.3303 | -103.3669
26.3341 | -103.3632
26.3378 | -103.3587
26.3416 | -103.3537
26.3454 | -103.3482
26.3491 | -103.3414
26.3529 | -103.336
26.3566 | -103.332
26.3604 | -103.322
26.3641 | -103.3116
26.3679 | -103.3018
26.3716 | -103.2936
26.3754 | -103.2834
26.3791 | -103.2735
26.3829 | -103.2641
26.3866 | -103.2554
26.3904 | -103.2474
26.3941 | -103.2391
26.3979 | -103.2313
26.4016 | -103.2238
26.4054 | -103.2126
26.4091 | -103.2014
26.4129 | -103.1896
26.4166 | -103.1788
26.4204 | -103.1697
26.4241 | -103.1604
26.4279 | -103.1514
26.4316 | -103.1424
26.4354 | -103.134
26.4391 | -103.1255
26.4429 | -103.1166
26.4466 | -103.1097
26.4504 | -103.1018
26.4541 | -103.0945
26.4579 | -103.0878
26.4616 | -103.0807
26.4654 | -103.0745
26.4691 | -103.0677
26.4729 | -103.0613
26.4766 | -103.0557
26.4804 | -103.0513
26.4841 | -103.0456
26.4879 | -103.0379
26.4916 | -103.0298
26.4954 | -103.0214
26.4991 | -103.0142
26.5029 | -103.007
26.5066 | -103.0005
26.5104 | -102.9941
26.5141 | -102.9742
26.5179 | -102.9587
26.5216 | -102.9451
26.5254 | -102.9336
26.5291 | -102.9216
26.5329 | -102.9099
26.5366 | -102.8974
26.5404 | -102.8847
26.5441 | -102.8727
26.5479 | -102.8622
26.5516 | -102.8532
26.5554 | -102.8457
26.5591 | -102.8366
26.5629 | -102.8235
26.5666 | -102.8111
26.5704 | -102.7994
26.5741 | -102.7862
26.5779 | -102.7735
26.5816 | -102.7606
26.5854 | -102.7479
26.5892 | -102.7335
26.5929 | -102.7195
26.5967 | -102.7054
26.6004 | -102.6919
26.6042 | -102.6787
26.6079 | -102.6654
26.6117 | -102.6534
26.6154 | -102.6408
26.6192 | -102.6275
26.6229 | -102.6145
26.6267 | -102.6016
26.6304 | -102.5877
26.6342 | -102.5729
26.6379 | -102.5593
26.6417 | -102.5482
26.6454 | -102.5368
26.6492 | -102.5251
26.6529 | -102.5162
26.6567 | -102.5064
26.6604 | -102.4941
26.6642 | -102.483
26.6679 | -102.4742
26.6717 | -102.4647
26.6754 | -102.4544
26.6792 | -102.4418
26.6829 | -102.4313
26.6867 | -102.4235
26.6904 | -102.4172
26.6942 | -102.4114
26.6979 | -102.4065
26.7017 | -102.4034
26.7054 | -102.3977
26.7092 | -102.3911
26.7129 | -102.3853
26.7167 | -102.381
26.7204 | -102.3773
26.7242 | -102.3751
26.7279 | -102.3732
26.7317 | -102.3709
26.7354 | -102.369
26.7392 | -102.3666
26.7429 | -102.3651
26.7467 | -102.363
26.7504 | -102.3604
26.7542 | -102.3567
26.7579 | -102.3536
26.7617 | -102.3515
26.7654 | -102.3494
26.7692 | -102.3476
26.7729 | -102.3436
26.7767 | -102.3378
26.7804 | -102.3324
26.7842 | -102.3288
26.7879 | -102.3254
26.7917 | -102.3211
26.7954 | -102.316
26.7992 | -102.3091
26.8029 | -102.3029
26.8067 | -102.2969
26.8104 | -102.2901
26.8142 | -102.2843
26.8179 | -102.2755
26.8217 | -102.2657
26.8254 | -102.2556
26.8292 | -102.2465
26.833 | -102.2376
26.8367 | -102.2284
26.8405 | -102.2182
26.8442 | -102.2085
26.848 | -102.1969
26.8517 | -102.1832
26.8555 | -102.1692
26.8592 | -102.157
26.863 | -102.1455
26.8667 | -102.1338
26.8705 | -102.1224
26.8742 | -102.1084
26.878 | -102.0932
26.8817 | -102.0798
26.8855 | -102.0658
26.8892 | -102.0533
26.893 | -102.0442
26.8967 | -102.0327
26.9005 | -102.0198
26.9042 | -102.0054
26.908 | -101.9902
26.9117 | -101.9739
26.9155 | -101.9577
26.9192 | -101.94
26.923 | -101.9214
26.9267 | -101.901
26.9305 | -101.8824
26.9342 | -101.864
26.938 | -101.8459
26.9417 | -101.8298
26.9455 | -101.8112
26.9492 | -101.792
26.953 | -101.7758
26.9567 | -101.7586
26.9605 | -101.7424
26.9642 | -101.7271
26.968 | -101.7123
26.9717 | -101.6968
26.9755 | -101.6815
26.9792 | -101.6654
26.983 | -101.6479
26.9867 | -101.6339
26.9905 | -101.6174
26.9942 | -101.6007
26.998 | -101.5824
27.0017 | -101.566
27.0055 | -101.5533
27.0092 | -101.5423
27.013 | -101.5309
27.0167 | -101.5203
27.0205 | -101.5063
27.0242 | -101.4894
27.028 | -101.4775
27.0317 | -101.4663
27.0355 | -101.4555
27.0392 | -101.4458
27.043 | -101.434
27.0467 | -101.4206
27.0505 | -101.4087
27.0542 | -101.3979
27.058 | -101.3882
27.0617 | -101.3787
27.0655 | -101.3694
27.0692 | -101.3598
27.073 | -101.3512
27.0768 | -101.3411
27.0805 | -101.3297
27.0843 | -101.3144
27.088 | -101.3005
27.0918 | -101.2885
27.0955 | -101.2758
27.0993 | -101.2638
27.103 | -101.2498
27.1068 | -101.2356
27.1105 | -101.2218
27.1143 | -101.2095
27.118 | -101.1993
27.1218 | -101.1889
27.1255 | -101.1769
27.1293 | -101.1651
27.133 | -101.1541
27.1368 | -101.1437
27.1405 | -101.1315
27.1443 | -101.1163
27.148 | -101.1018
27.1518 | -101.0896
27.1555 | -101.0798
27.1593 | -101.0699
27.163 | -101.0602
27.1668 | -101.0494
27.1705 | -101.0379
27.1743 | -101.0273
27.178 | -101.0168
27.1818 | -101.0068
27.1855 | -100.9966
27.1893 | -100.9878
27.193 | -100.9799
27.1968 | -100.9707
27.2005 | -100.9632
27.2043 | -100.9573
27.208 | -100.9521
27.2118 | -100.9451
27.2155 | -100.9376
27.2193 | -100.9312
27.223 | -100.9265
27.2268 | -100.9226
27.2305 | -100.9203
27.2343 | -100.9191
27.238 | -100.9203
27.2418 | -100.9209
27.2455 | -100.9206
27.2493 | -100.921
27.253 | -100.9219
27.2568 | -100.9248
27.2605 | -100.9282
27.2643 | -100.9309
27.268 | -100.9354
27.2718 | -100.9419
27.2755 | -100.9482
27.2793 | -100.9563
27.283 | -100.9645
27.2868 | -100.9733
27.2905 | -100.9839
27.2943 | -100.9961
27.298 | -101.0101
27.3018 | -101.0242
27.3055 | -101.0395
27.3093 | -101.0575
27.313 | -101.0735
27.3168 | -101.0913
27.3205 | -101.1096
27.3243 | -101.1276
27.3281 | -101.1497
27.3318 | -101.173
27.3356 | -101.1941
27.3393 | -101.2164
27.3431 | -101.2386
27.3468 | -101.2611
27.3506 | -101.2849
27.3543 | -101.3093
27.3581 | -101.3326
27.3618 | -101.3562
27.3656 | -101.3833
27.3693 | -101.4067
27.3731 | -101.4331
27.3768 | -101.4612
27.3806 | -101.4895
27.3843 | -101.515
27.3881 | -101.5378
27.3918 | -101.5606
27.3956 | -101.584
27.3993 | -101.6061
27.4031 | -101.6309
27.4068 | -101.6584
27.4106 | -101.68
27.4143 | -101.6996
27.4181 | -101.7197
27.4218 | -101.738
27.4256 | -101.7564
27.4293 | -101.7746
27.4331 | -101.7915
27.4368 | -101.8095
27.4406 | -101.8247
27.4443 | -101.8392
27.4481 | -101.8546
27.4518 | -101.8685
27.4556 | -101.8807
27.4593 | -101.8912
27.4631 | -101.9033
27.4668 | -101.9163
27.4706 | -101.9337
27.4743 | -101.9498
27.4781 | -101.9603
27.4818 | -101.9694
27.4856 | -101.9783
27.4893 | -101.9873
27.4931 | -101.9956
27.4968 | -102.0048
27.5006 | -102.0104
27.5043 | -102.011
27.5081 | -102.0125
27.5118 | -102.0151
27.5156 | -102.0176
27.5193 | -102.0206
27.5231 | -102.0234
27.5268 | -102.0277
27.5306 | -102.0332
27.5343 | -102.0362
27.5381 | -102.0331
27.5418 | -102.0278
27.5456 | -102.0237
27.5493 | -102.0184
27.5531 | -102.0096
27.5568 | -102.0009
27.5606 | -101.9939
27.5643 | -101.9878
27.5681 | -101.9826
27.5719 | -101.9779
27.5756 | -101.9716
27.5794 | -101.9645
27.5831 | -101.9579
27.5869 | -101.9527
27.5906 | -101.9457
27.5944 | -101.9311
27.5981 | -101.9207
27.6019 | -101.9118
27.6056 | -101.905
27.6094 | -101.8969
27.6131 | -101.8869
27.6169 | -101.8765
27.6206 | -101.8644
27.6244 | -101.8531
27.6281 | -101.8401
27.6319 | -101.8282
27.6356 | -101.8164
27.6394 | -101.8046
27.6431 | -101.7933
27.6469 | -101.782
27.6506 | -101.7703
27.6544 | -101.7579
27.6581 | -101.7451
27.6619 | -101.7316
27.6656 | -101.7194
27.6694 | -101.7082
27.6731 | -101.6968
27.6769 | -101.6855
27.6806 | -101.6754
27.6844 | -101.6662
27.6881 | -101.6555
27.6919 | -101.6468
27.6956 | -101.6375
27.6994 | -101.628
27.7031 | -101.6211
27.7069 | -101.6123
27.7106 | -101.6054
27.7144 | -101.6008
27.7181 | -101.5958
27.7219 | -101.5913
27.7256 | -101.5881
27.7294 | -101.5854
27.7331 | -101.582
27.7369 | -101.5806
27.7406 | -101.5801
27.7444 | -101.5799
27.7481 | -101.5826
27.7519 | -101.5849
27.7556 | -101.5865
27.7594 | -101.5895
27.7631 | -101.5913
27.7669 | -101.5927
27.7706 | -101.5967
27.7744 | -101.602
27.7781 | -101.6085
27.7819 | -101.6152
27.7856 | -101.6216
27.7894 | -101.6288
27.7931 | -101.6368
27.7969 | -101.6446
27.8006 | -101.6521
27.8044 | -101.6612
27.8081 | -101.6704
27.8119 | -101.681
27.8157 | -101.6924
27.8194 | -101.7052
27.8232 | -101.7178
27.8269 | -101.7301
27.8307 | -101.7418
27.8344 | -101.7522
27.8382 | -101.7626
27.8419 | -101.7715
27.8457 | -101.7818
27.8494 | -101.7932
27.8532 | -101.8031
27.8569 | -101.8105
27.8607 | -101.8183
27.8644 | -101.8286
27.8682 | -101.84
27.8719 | -101.8512
27.8757 | -101.8609
27.8794 | -101.871
27.8832 | -101.8809
27.8869 | -101.8881
27.8907 | -101.8953
27.8944 | -101.9057
27.8982 | -101.914
27.9019 | -101.9219
27.9057 | -101.9284
27.9094 | -101.9338
27.9132 | -101.939
27.9169 | -101.9441
27.9207 | -101.9477
27.9244 | -101.9501
27.9282 | -101.9535
27.9319 | -101.9571
27.9357 | -101.9612
27.9394 | -101.9653
27.9432 | -101.9713
27.9469 | -101.9756
27.9507 | -101.976
27.9544 | -101.9742
27.9582 | -101.9723
27.9619 | -101.9669
27.9657 | -101.9608
27.9694 | -101.9562
27.9732 | -101.9529
27.9769 | -101.944
27.9807 | -101.935
27.9844 | -101.9303
27.9882 | -101.925
27.9919 | -101.9188
27.9957 | -101.9105
27.9994 | -101.9008
28.0032 | -101.8876