## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 251 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 269
Hood | 02-92B/P | 196
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1015.4 | hPa |  | PRC PRO 01
Humidity | 54.5 | %/Hr
Internal temperature of calibration bench | 23.31 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1022 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.3 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.02 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.04 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3065 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 18/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 22/06/2018
Technical Manager | Anthony HUE | Approval |  | 28/06/2018
CEO | Robert GLEMOT | Approval |  | 28/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3573 | 57.4684
0.012209 | -35.8839 | 38.3567
0.018313 | -34.8342 | 28.442
0.024417 | -34.3884 | 22.0581
0.030521 | -34.2892 | 18.2939
0.036626 | -34.1496 | 15.1933
0.04273 | -34.0663 | 12.8701
0.048834 | -34.0144 | 11.3039
0.054938 | -33.9415 | 10.11
0.061043 | -33.8641 | 8.9659
0.067147 | -33.8544 | 8.0897
0.073251 | -33.8566 | 7.3328
0.079355 | -33.8764 | 6.7754
0.08546 | -33.8961 | 6.3168
0.091564 | -33.8933 | 5.941
0.097668 | -33.8587 | 5.6898
0.10377 | -33.8101 | 5.3456
0.10988 | -33.7785 | 4.9835
0.11598 | -33.7694 | 4.6223
0.12209 | -33.7868 | 4.342
0.12819 | -33.8208 | 4.1078
0.13429 | -33.8539 | 3.8656
0.1404 | -33.8713 | 3.5892
0.1465 | -33.8872 | 3.3285
0.15261 | -33.8942 | 3.0873
0.15871 | -33.8903 | 2.8913
0.16482 | -33.8755 | 2.7351
0.17092 | -33.8508 | 2.6775
0.17702 | -33.8319 | 2.5827
0.18313 | -33.8146 | 2.5018
0.18923 | -33.8022 | 2.3958
0.19534 | -33.7955 | 2.2912
0.20144 | -33.7947 | 2.1756
0.20754 | -33.7975 | 2.1005
0.21365 | -33.8001 | 2.0247
0.21975 | -33.8046 | 1.9598
0.22586 | -33.8128 | 1.8948
0.23196 | -33.8216 | 1.8228
0.23807 | -33.8305 | 1.7459
0.24417 | -33.8328 | 1.6528
0.25027 | -33.8363 | 1.5688
0.25638 | -33.8363 | 1.4767
0.26248 | -33.8344 | 1.3894
0.26859 | -33.832 | 1.3057
0.27469 | -33.8312 | 1.2182
0.2808 | -33.8373 | 1.1528
0.2869 | -33.8419 | 1.0833
0.293 | -33.8476 | 1.0251
0.29911 | -33.8542 | 0.98604
0.30521 | -33.8605 | 0.94862
0.31132 | -33.8648 | 0.93507
0.31742 | -33.8694 | 0.91303
0.32353 | -33.8731 | 0.89543
0.32963 | -33.8762 | 0.88124
0.33573 | -33.8783 | 0.86953
0.34184 | -33.881 | 0.8544
0.34794 | -33.8843 | 0.82058
0.35405 | -33.8878 | 0.79177
0.36015 | -33.8915 | 0.75937
0.36626 | -33.894 | 0.72328
0.37236 | -33.8965 | 0.67389
0.37846 | -33.8988 | 0.61648
0.38457 | -33.8978 | 0.5564
0.39067 | -33.8984 | 0.50583
0.39678 | -33.8967 | 0.4487
0.40288 | -33.8943 | 0.39456
0.40899 | -33.8904 | 0.34161
0.41509 | -33.8857 | 0.28838
0.42119 | -33.8806 | 0.24014
0.4273 | -33.8768 | 0.19024
0.4334 | -33.8733 | 0.1395
0.43951 | -33.8702 | 0.089598
0.44561 | -33.8686 | 0.043064
0.45172 | -33.8703 | 0.0048865
0.45782 | -33.8701 | -0.037811
0.46392 | -33.8714 | -0.077087
0.47003 | -33.8732 | -0.1126
0.47613 | -33.8753 | -0.14505
0.48224 | -33.8776 | -0.176
0.48834 | -33.8802 | -0.19777
0.49445 | -33.8823 | -0.22287
0.50055 | -33.8849 | -0.24552
0.50665 | -33.8867 | -0.2641
0.51276 | -33.888 | -0.2842
0.51886 | -33.8892 | -0.30176
0.52497 | -33.8883 | -0.30814
0.53107 | -33.8887 | -0.31842
0.53717 | -33.8891 | -0.32969
0.54328 | -33.889 | -0.33472
0.54938 | -33.889 | -0.34265
0.55549 | -33.8886 | -0.35145
0.56159 | -33.8886 | -0.36214
0.5677 | -33.8882 | -0.37229
0.5738 | -33.8878 | -0.38535
0.5799 | -33.888 | -0.39715
0.58601 | -33.8885 | -0.4133
0.59211 | -33.8897 | -0.43547
0.59822 | -33.8904 | -0.45176
0.60432 | -33.8916 | -0.47318
0.61043 | -33.8927 | -0.49411
0.61653 | -33.894 | -0.51542
0.62263 | -33.8951 | -0.53826
0.62874 | -33.8966 | -0.55826
0.63484 | -33.8978 | -0.57908
0.64095 | -33.8991 | -0.59978
0.64705 | -33.9003 | -0.61945
0.65316 | -33.9013 | -0.63872
0.65926 | -33.9024 | -0.66121
0.66536 | -33.9018 | -0.67949
0.67147 | -33.9021 | -0.69558
0.67757 | -33.9026 | -0.71331
0.68368 | -33.9024 | -0.73333
0.68978 | -33.9025 | -0.75123
0.69589 | -33.9021 | -0.76949
0.70199 | -33.9012 | -0.79154
0.70809 | -33.9006 | -0.81172
0.7142 | -33.9001 | -0.83102
0.7203 | -33.8992 | -0.85063
0.72641 | -33.8983 | -0.87157
0.73251 | -33.898 | -0.89241
0.73862 | -33.898 | -0.91684
0.74472 | -33.8983 | -0.93619
0.75082 | -33.8978 | -0.95734
0.75693 | -33.8979 | -0.981
0.76303 | -33.8978 | -1.0036
0.76914 | -33.8984 | -1.0308
0.77524 | -33.8987 | -1.0559
0.78135 | -33.8991 | -1.0794
0.78745 | -33.8998 | -1.1043
0.79355 | -33.9004 | -1.1291
0.79966 | -33.9014 | -1.1551
0.80576 | -33.9027 | -1.1857
0.81187 | -33.9031 | -1.2155
0.81797 | -33.9038 | -1.2466
0.82408 | -33.9046 | -1.2741
0.83018 | -33.9052 | -1.3056
0.83628 | -33.9062 | -1.3364
0.84239 | -33.9061 | -1.3649
0.84849 | -33.9062 | -1.3933
0.8546 | -33.9067 | -1.4213
0.8607 | -33.9068 | -1.451
0.86681 | -33.907 | -1.4797
0.87291 | -33.9066 | -1.5056
0.87901 | -33.9062 | -1.5282
0.88512 | -33.9061 | -1.5514
0.89122 | -33.906 | -1.575
0.89733 | -33.9058 | -1.5943
0.90343 | -33.9055 | -1.6157
0.90953 | -33.9057 | -1.6305
0.91564 | -33.9052 | -1.6465
0.92174 | -33.9048 | -1.6606
0.92785 | -33.9044 | -1.674
0.93395 | -33.9044 | -1.6856
0.94006 | -33.9044 | -1.699
0.94616 | -33.9042 | -1.7033
0.95226 | -33.9043 | -1.7145
0.95837 | -33.9045 | -1.726
0.96447 | -33.9049 | -1.736
0.97058 | -33.9049 | -1.7456
0.97668 | -33.9047 | -1.7533
0.98279 | -33.9047 | -1.7565
0.98889 | -33.905 | -1.7609
0.99499 | -33.905 | -1.7692
1.0011 | -33.9053 | -1.7736
1.0072 | -33.9052 | -1.7769
1.0133 | -33.9051 | -1.7803
1.0194 | -33.9051 | -1.7854
1.0255 | -33.9048 | -1.792
1.0316 | -33.9041 | -1.7956
1.0377 | -33.9036 | -1.8012
1.0438 | -33.9028 | -1.8053
1.0499 | -33.9012 | -1.8114
1.056 | -33.9 | -1.8159
1.0621 | -33.8988 | -1.8214
1.0682 | -33.8983 | -1.8259
1.0743 | -33.8971 | -1.8336
1.0805 | -33.8965 | -1.8401
1.0866 | -33.8948 | -1.8507
1.0927 | -33.8933 | -1.8579
1.0988 | -33.8925 | -1.8683
1.1049 | -33.891 | -1.8756
1.111 | -33.8893 | -1.8866
1.1171 | -33.8879 | -1.8957
1.1232 | -33.8865 | -1.9099
1.1293 | -33.8853 | -1.9217
1.1354 | -33.8843 | -1.934
1.1415 | -33.8831 | -1.948
1.1476 | -33.8816 | -1.9619
1.1537 | -33.8802 | -1.9751
1.1598 | -33.8801 | -1.9935
1.1659 | -33.8795 | -2.0054
1.172 | -33.8788 | -2.0186
1.1781 | -33.878 | -2.0326
1.1842 | -33.8774 | -2.0482
1.1903 | -33.8782 | -2.0633
1.1964 | -33.878 | -2.0777
1.2025 | -33.878 | -2.0929
1.2086 | -33.8784 | -2.1086
1.2147 | -33.8787 | -2.1235
1.2209 | -33.8794 | -2.1369
1.227 | -33.8814 | -2.1518
1.2331 | -33.8824 | -2.1661
1.2392 | -33.8834 | -2.1819
1.2453 | -33.8846 | -2.1937
1.2514 | -33.886 | -2.2057
1.2575 | -33.8876 | -2.2196
1.2636 | -33.89 | -2.2309
1.2697 | -33.8918 | -2.2423
1.2758 | -33.8935 | -2.2564
1.2819 | -33.8953 | -2.2698
1.288 | -33.897 | -2.2835
1.2941 | -33.8992 | -2.2939
1.3002 | -33.9012 | -2.3019
1.3063 | -33.903 | -2.3117
1.3124 | -33.9049 | -2.3215
1.3185 | -33.9069 | -2.332
1.3246 | -33.9088 | -2.3425
1.3307 | -33.911 | -2.3518
1.3368 | -33.9129 | -2.3591
1.3429 | -33.9147 | -2.3676
1.349 | -33.9162 | -2.3759
1.3551 | -33.918 | -2.3866
1.3613 | -33.9196 | -2.3942
1.3674 | -33.9209 | -2.4025
1.3735 | -33.9223 | -2.4104
1.3796 | -33.9236 | -2.4175
1.3857 | -33.9246 | -2.4242
1.3918 | -33.9257 | -2.4338
1.3979 | -33.9271 | -2.4421
1.404 | -33.9278 | -2.452
1.4101 | -33.9288 | -2.4608
1.4162 | -33.9296 | -2.4692
1.4223 | -33.9305 | -2.4794
1.4284 | -33.9313 | -2.4875
1.4345 | -33.9319 | -2.4977
1.4406 | -33.9324 | -2.5089
1.4467 | -33.9331 | -2.5182
1.4528 | -33.9338 | -2.5273
1.4589 | -33.9347 | -2.5378
1.465 | -33.9349 | -2.5453
1.4711 | -33.9359 | -2.5542
1.4772 | -33.9364 | -2.5664
1.4833 | -33.9371 | -2.5786
1.4894 | -33.9378 | -2.59
1.4955 | -33.9386 | -2.6015
1.5016 | -33.9393 | -2.6135
1.5078 | -33.9397 | -2.6282
1.5139 | -33.9404 | -2.6409
1.52 | -33.9409 | -2.6541
1.5261 | -33.9414 | -2.6657
1.5322 | -33.9416 | -2.6784
1.5383 | -33.9419 | -2.6905
1.5444 | -33.9425 | -2.7039
1.5505 | -33.943 | -2.7172
1.5566 | -33.9435 | -2.73
1.5627 | -33.9441 | -2.7435
1.5688 | -33.9449 | -2.7567
1.5749 | -33.9452 | -2.769
1.581 | -33.9456 | -2.7858
1.5871 | -33.9461 | -2.8015
1.5932 | -33.9466 | -2.8144
1.5993 | -33.9469 | -2.8313
1.6054 | -33.9473 | -2.8464
1.6115 | -33.9477 | -2.8634
1.6176 | -33.948 | -2.8799
1.6237 | -33.9483 | -2.8944
1.6298 | -33.9486 | -2.9099
1.6359 | -33.949 | -2.9244
1.642 | -33.9495 | -2.9389
1.6482 | -33.9499 | -2.9561
1.6543 | -33.9504 | -2.973
1.6604 | -33.9508 | -2.9877
1.6665 | -33.9509 | -3.0026
1.6726 | -33.9512 | -3.019
1.6787 | -33.9517 | -3.0361
1.6848 | -33.9516 | -3.0526
1.6909 | -33.952 | -3.069
1.697 | -33.9522 | -3.0853
1.7031 | -33.9523 | -3.1011
1.7092 | -33.9523 | -3.116
1.7153 | -33.9525 | -3.1323
1.7214 | -33.952 | -3.1478
1.7275 | -33.9519 | -3.1637
1.7336 | -33.9516 | -3.18
1.7397 | -33.9517 | -3.1954
1.7458 | -33.9517 | -3.2121
1.7519 | -33.9515 | -3.2284
1.758 | -33.9515 | -3.2434
1.7641 | -33.9516 | -3.2579
1.7702 | -33.9513 | -3.2738
1.7763 | -33.9513 | -3.2875
1.7824 | -33.9512 | -3.3038
1.7885 | -33.951 | -3.318
1.7947 | -33.9511 | -3.3329
1.8008 | -33.951 | -3.3458
1.8069 | -33.9509 | -3.3595
1.813 | -33.9509 | -3.3738
1.8191 | -33.9506 | -3.3881
1.8252 | -33.9504 | -3.4022
1.8313 | -33.9504 | -3.4166
1.8374 | -33.9504 | -3.4302
1.8435 | -33.9505 | -3.4424
1.8496 | -33.9502 | -3.457
1.8557 | -33.9501 | -3.4719
1.8618 | -33.9501 | -3.4837
1.8679 | -33.9501 | -3.4968
1.874 | -33.9501 | -3.5099
1.8801 | -33.95 | -3.523
1.8862 | -33.9501 | -3.5358
1.8923 | -33.95 | -3.548
1.8984 | -33.9497 | -3.5587
1.9045 | -33.9496 | -3.5697
1.9106 | -33.9493 | -3.5816
1.9167 | -33.949 | -3.5929
1.9228 | -33.949 | -3.6037
1.9289 | -33.9487 | -3.6126
1.9351 | -33.9487 | -3.6216
1.9412 | -33.9488 | -3.6318
1.9473 | -33.9489 | -3.6438
1.9534 | -33.9488 | -3.6549
1.9595 | -33.9486 | -3.6664
1.9656 | -33.9481 | -3.6772
1.9717 | -33.9479 | -3.6877
1.9778 | -33.9479 | -3.6992
1.9839 | -33.9479 | -3.7106
1.99 | -33.948 | -3.7204
1.9961 | -33.9477 | -3.7308
2.0022 | -33.9476 | -3.7407
2.0083 | -33.9473 | -3.7506
2.0144 | -33.9472 | -3.7605
2.0205 | -33.9473 | -3.7723
2.0266 | -33.9474 | -3.7839
2.0327 | -33.9474 | -3.795
2.0388 | -33.9477 | -3.8089
2.0449 | -33.9477 | -3.8191
2.051 | -33.9478 | -3.8291
2.0571 | -33.9477 | -3.8404
2.0632 | -33.9475 | -3.851
2.0693 | -33.9475 | -3.8621
2.0754 | -33.9478 | -3.8739
2.0816 | -33.9479 | -3.886
2.0877 | -33.9478 | -3.897
2.0938 | -33.9481 | -3.9094
2.0999 | -33.948 | -3.9211
2.106 | -33.9483 | -3.933
2.1121 | -33.9482 | -3.9453
2.1182 | -33.9484 | -3.9553
2.1243 | -33.9486 | -3.9658
2.1304 | -33.9487 | -3.9761
2.1365 | -33.949 | -3.9876
2.1426 | -33.9491 | -4.001
2.1487 | -33.9496 | -4.0127
2.1548 | -33.95 | -4.0258
2.1609 | -33.9506 | -4.0378
2.167 | -33.9511 | -4.0504
2.1731 | -33.9514 | -4.0624
2.1792 | -33.9521 | -4.0736
2.1853 | -33.9526 | -4.0865
2.1914 | -33.9531 | -4.0987
2.1975 | -33.9537 | -4.1109
2.2036 | -33.9541 | -4.1242
2.2097 | -33.9546 | -4.1362
2.2158 | -33.9555 | -4.1465
2.222 | -33.9561 | -4.1566
2.2281 | -33.9564 | -4.1674
2.2342 | -33.9569 | -4.1805
2.2403 | -33.9574 | -4.1906
2.2464 | -33.9582 | -4.2017
2.2525 | -33.9586 | -4.2125
2.2586 | -33.9595 | -4.2223
2.2647 | -33.9604 | -4.2316
2.2708 | -33.961 | -4.2433
2.2769 | -33.9619 | -4.2548
2.283 | -33.9623 | -4.2642
2.2891 | -33.9629 | -4.2746
2.2952 | -33.9635 | -4.2844
2.3013 | -33.9641 | -4.2946
2.3074 | -33.9649 | -4.3042
2.3135 | -33.9653 | -4.3156
2.3196 | -33.966 | -4.3241
2.3257 | -33.9663 | -4.3347
2.3318 | -33.9667 | -4.3437
2.3379 | -33.9672 | -4.3552
2.344 | -33.9675 | -4.3661
2.3501 | -33.968 | -4.3768
2.3562 | -33.9683 | -4.3851
2.3623 | -33.9686 | -4.3938
2.3685 | -33.9688 | -4.4028
2.3746 | -33.9691 | -4.4133
2.3807 | -33.9696 | -4.4223
2.3868 | -33.9698 | -4.4313
2.3929 | -33.9702 | -4.4406
2.399 | -33.9705 | -4.4486
2.4051 | -33.971 | -4.4581
2.4112 | -33.9713 | -4.4671
2.4173 | -33.9716 | -4.476
2.4234 | -33.9716 | -4.4845
2.4295 | -33.9719 | -4.4931
2.4356 | -33.972 | -4.5025
2.4417 | -33.9721 | -4.5125
2.4478 | -33.9722 | -4.5221
2.4539 | -33.9725 | -4.53
2.46 | -33.9722 | -4.5397
2.4661 | -33.9722 | -4.549
2.4722 | -33.9725 | -4.5577
2.4783 | -33.9726 | -4.5661
2.4844 | -33.9726 | -4.5747
2.4905 | -33.9725 | -4.5839
2.4966 | -33.9722 | -4.5928
2.5027 | -33.9722 | -4.6008
2.5089 | -33.9721 | -4.6093
2.515 | -33.9719 | -4.6186
2.5211 | -33.9719 | -4.6269
2.5272 | -33.9715 | -4.636
2.5333 | -33.9713 | -4.6456
2.5394 | -33.9713 | -4.6546
2.5455 | -33.9711 | -4.664
2.5516 | -33.9711 | -4.6728
2.5577 | -33.971 | -4.6828
2.5638 | -33.9708 | -4.6932
2.5699 | -33.9705 | -4.7035
2.576 | -33.9703 | -4.714
2.5821 | -33.9701 | -4.724
2.5882 | -33.9699 | -4.7336
2.5943 | -33.9698 | -4.7434
2.6004 | -33.9696 | -4.7535
2.6065 | -33.9694 | -4.7619
2.6126 | -33.9693 | -4.771
2.6187 | -33.9692 | -4.7798
2.6248 | -33.9691 | -4.7889
2.6309 | -33.9689 | -4.8002
2.637 | -33.969 | -4.8115
2.6431 | -33.9689 | -4.821
2.6492 | -33.9688 | -4.8311
2.6554 | -33.9688 | -4.8414
2.6615 | -33.9686 | -4.8513
2.6676 | -33.9685 | -4.8619
2.6737 | -33.9687 | -4.8727
2.6798 | -33.9687 | -4.8827
2.6859 | -33.9687 | -4.8933
2.692 | -33.9686 | -4.9031
2.6981 | -33.9685 | -4.9143
2.7042 | -33.9686 | -4.9265
2.7103 | -33.9686 | -4.9366
2.7164 | -33.9688 | -4.9474
2.7225 | -33.9689 | -4.9588
2.7286 | -33.969 | -4.9701
2.7347 | -33.969 | -4.9801
2.7408 | -33.9693 | -4.9893
2.7469 | -33.9693 | -5.0008
2.753 | -33.9696 | -5.0115
2.7591 | -33.9697 | -5.0236
2.7652 | -33.9697 | -5.0353
2.7713 | -33.9697 | -5.0452
2.7774 | -33.9702 | -5.0567
2.7835 | -33.9703 | -5.0674
2.7896 | -33.9704 | -5.0768
2.7958 | -33.9708 | -5.0871
2.8019 | -33.9714 | -5.097
2.808 | -33.9717 | -5.1089
2.8141 | -33.9721 | -5.122
2.8202 | -33.9725 | -5.1343
2.8263 | -33.9729 | -5.1453
2.8324 | -33.9732 | -5.1578
2.8385 | -33.9736 | -5.17
2.8446 | -33.9742 | -5.1818
2.8507 | -33.9746 | -5.194
2.8568 | -33.975 | -5.2052
2.8629 | -33.9755 | -5.2166
2.869 | -33.9759 | -5.2281
2.8751 | -33.9764 | -5.2401
2.8812 | -33.977 | -5.2535
2.8873 | -33.9774 | -5.265
2.8934 | -33.9777 | -5.277
2.8995 | -33.9781 | -5.2891
2.9056 | -33.9784 | -5.2999
2.9117 | -33.9789 | -5.3116
2.9178 | -33.9796 | -5.3244
2.9239 | -33.98 | -5.3378
2.93 | -33.9805 | -5.3515
2.9361 | -33.9811 | -5.363
2.9423 | -33.9817 | -5.3751
2.9484 | -33.9822 | -5.3887
2.9545 | -33.9829 | -5.402
2.9606 | -33.9834 | -5.4155
2.9667 | -33.984 | -5.4275
2.9728 | -33.9844 | -5.4398
2.9789 | -33.9849 | -5.4527
2.985 | -33.9853 | -5.4666
2.9911 | -33.9858 | -5.4793
2.9972 | -33.9864 | -5.4921
3.0033 | -33.987 | -5.5053
3.0094 | -33.9875 | -5.5185
3.0155 | -33.9881 | -5.5312
3.0216 | -33.9886 | -5.5447
3.0277 | -33.9891 | -5.5586
3.0338 | -33.9897 | -5.5723
3.0399 | -33.9903 | -5.5862
3.046 | -33.9906 | -5.5994
3.0521 | -33.9912 | -5.6119
3.0582 | -33.9915 | -5.626
3.0643 | -33.9921 | -5.6391
3.0704 | -33.9923 | -5.6527
3.0765 | -33.9927 | -5.665
3.0827 | -33.993 | -5.6782
3.0888 | -33.9934 | -5.6912
3.0949 | -33.9937 | -5.7043
3.101 | -33.9942 | -5.7186
3.1071 | -33.9945 | -5.7328
3.1132 | -33.995 | -5.7471
3.1193 | -33.9954 | -5.7602
3.1254 | -33.9958 | -5.7725
3.1315 | -33.9963 | -5.787
3.1376 | -33.9965 | -5.8
3.1437 | -33.9969 | -5.8137
3.1498 | -33.9972 | -5.828
3.1559 | -33.9976 | -5.8414
3.162 | -33.9976 | -5.8557
3.1681 | -33.9977 | -5.87
3.1742 | -33.9981 | -5.8839
3.1803 | -33.9984 | -5.8972
3.1864 | -33.9988 | -5.9096
3.1925 | -33.9991 | -5.9234
3.1986 | -33.9992 | -5.9363
3.2047 | -33.9993 | -5.9505
3.2108 | -33.9996 | -5.9623
3.2169 | -34 | -5.9762
3.223 | -33.9999 | -5.9897
3.2292 | -33.9999 | -6.0038
3.2353 | -33.9998 | -6.0133
3.2414 | -34 | -6.0264
3.2475 | -33.9999 | -6.0375
3.2536 | -34.0002 | -6.0514
3.2597 | -34.0005 | -6.0627
3.2658 | -34.0006 | -6.0735
3.2719 | -34.0008 | -6.0846
3.278 | -34.0008 | -6.0969
3.2841 | -34.0009 | -6.1089
3.2902 | -34.001 | -6.1215
3.2963 | -34.0012 | -6.1333
3.3024 | -34.001 | -6.1435
3.3085 | -34.0013 | -6.1562
3.3146 | -34.0012 | -6.1672
3.3207 | -34.0014 | -6.1794
3.3268 | -34.0015 | -6.1917
3.3329 | -34.0016 | -6.2015
3.339 | -34.0017 | -6.2112
3.3451 | -34.0017 | -6.2227
3.3512 | -34.0018 | -6.2322
3.3573 | -34.002 | -6.2436
3.3634 | -34.0021 | -6.2552
3.3696 | -34.0021 | -6.2672
3.3757 | -34.002 | -6.277
3.3818 | -34.0024 | -6.2888
3.3879 | -34.0023 | -6.2992
3.394 | -34.0022 | -6.3098
3.4001 | -34.0023 | -6.3207
3.4062 | -34.0023 | -6.3309
3.4123 | -34.0023 | -6.3407
3.4184 | -34.0023 | -6.3518
3.4245 | -34.0024 | -6.3624
3.4306 | -34.0024 | -6.3721
3.4367 | -34.0025 | -6.3819
3.4428 | -34.0028 | -6.3911
3.4489 | -34.0027 | -6.4022
3.455 | -34.0026 | -6.4127
3.4611 | -34.0027 | -6.4239
3.4672 | -34.0026 | -6.4353
3.4733 | -34.0027 | -6.4454
3.4794 | -34.0027 | -6.456
3.4855 | -34.0029 | -6.4656
3.4916 | -34.003 | -6.4753
3.4977 | -34.003 | -6.4863
3.5038 | -34.0031 | -6.4963
3.5099 | -34.0032 | -6.5059
3.5161 | -34.0033 | -6.5154
3.5222 | -34.0034 | -6.5254
3.5283 | -34.0033 | -6.5355
3.5344 | -34.0034 | -6.5463
3.5405 | -34.0035 | -6.5562
3.5466 | -34.0036 | -6.566
3.5527 | -34.0037 | -6.5767
3.5588 | -34.0038 | -6.5865
3.5649 | -34.0038 | -6.5966
3.571 | -34.0039 | -6.6069
3.5771 | -34.004 | -6.6177
3.5832 | -34.0041 | -6.6272
3.5893 | -34.0043 | -6.6371
3.5954 | -34.0044 | -6.6477
3.6015 | -34.0046 | -6.658
3.6076 | -34.0048 | -6.6677
3.6137 | -34.0049 | -6.6775
3.6198 | -34.0052 | -6.6865
3.6259 | -34.0053 | -6.6973
3.632 | -34.0055 | -6.7079
3.6381 | -34.0057 | -6.7178
3.6442 | -34.0059 | -6.7283
3.6503 | -34.0062 | -6.7373
3.6565 | -34.0065 | -6.7491
3.6626 | -34.0065 | -6.7598
3.6687 | -34.0068 | -6.7705
3.6748 | -34.0072 | -6.7804
3.6809 | -34.0075 | -6.7909
3.687 | -34.0077 | -6.8021
3.6931 | -34.0081 | -6.812
3.6992 | -34.0083 | -6.8227
3.7053 | -34.0085 | -6.8328
3.7114 | -34.0087 | -6.8441
3.7175 | -34.0091 | -6.854
3.7236 | -34.0093 | -6.8628
3.7297 | -34.0094 | -6.8735
3.7358 | -34.0096 | -6.8852
3.7419 | -34.0098 | -6.8969
3.748 | -34.01 | -6.9076
3.7541 | -34.0101 | -6.9186
3.7602 | -34.0105 | -6.9295
3.7663 | -34.0107 | -6.9396
3.7724 | -34.011 | -6.9507
3.7785 | -34.0114 | -6.9611
3.7846 | -34.0116 | -6.9717
3.7907 | -34.012 | -6.9828
3.7969 | -34.0122 | -6.9943
3.803 | -34.0126 | -7.0043
3.8091 | -34.0129 | -7.0146
3.8152 | -34.0132 | -7.0259
3.8213 | -34.0134 | -7.0375
3.8274 | -34.0138 | -7.048
3.8335 | -34.0142 | -7.0596
3.8396 | -34.0146 | -7.0711
3.8457 | -34.015 | -7.0814
3.8518 | -34.0155 | -7.0923
3.8579 | -34.0158 | -7.1029
3.864 | -34.0162 | -7.1136
3.8701 | -34.0166 | -7.1247
3.8762 | -34.0169 | -7.1359
3.8823 | -34.0172 | -7.1465
3.8884 | -34.0176 | -7.1569
3.8945 | -34.0181 | -7.1683
3.9006 | -34.0185 | -7.1799
3.9067 | -34.0189 | -7.1918
3.9128 | -34.0193 | -7.2031
3.9189 | -34.0197 | -7.2138
3.925 | -34.0199 | -7.2247
3.9311 | -34.0203 | -7.2347
3.9372 | -34.0207 | -7.2463
3.9434 | -34.0211 | -7.2567
3.9495 | -34.0215 | -7.2684
3.9556 | -34.0219 | -7.2796
3.9617 | -34.0223 | -7.2905
3.9678 | -34.0227 | -7.3012
3.9739 | -34.0231 | -7.3125
3.98 | -34.0235 | -7.3243
3.9861 | -34.0238 | -7.3348
3.9922 | -34.024 | -7.3465
3.9983 | -34.0244 | -7.3566
4.0044 | -34.0247 | -7.3663
4.0105 | -34.0251 | -7.3772
4.0166 | -34.0256 | -7.3883
4.0227 | -34.026 | -7.3986
4.0288 | -34.0262 | -7.4101
4.0349 | -34.0266 | -7.42
4.041 | -34.0269 | -7.4309
4.0471 | -34.0271 | -7.4412
4.0532 | -34.0276 | -7.4524
4.0593 | -34.0279 | -7.464
4.0654 | -34.0282 | -7.4748
4.0715 | -34.0286 | -7.4862
4.0776 | -34.0289 | -7.4966
4.0838 | -34.0293 | -7.5081
4.0899 | -34.0296 | -7.5194
4.096 | -34.0299 | -7.5298
4.1021 | -34.0303 | -7.5406
4.1082 | -34.0306 | -7.5509
4.1143 | -34.031 | -7.5608
4.1204 | -34.0314 | -7.5716
4.1265 | -34.0318 | -7.582
4.1326 | -34.0322 | -7.5923
4.1387 | -34.0326 | -7.6036
4.1448 | -34.0329 | -7.6146
4.1509 | -34.0332 | -7.6253
4.157 | -34.0335 | -7.6366
4.1631 | -34.0338 | -7.6478
4.1692 | -34.0341 | -7.6586
4.1753 | -34.0345 | -7.6695
4.1814 | -34.0347 | -7.6793
4.1875 | -34.0351 | -7.6907
4.1936 | -34.0353 | -7.7021
4.1997 | -34.0357 | -7.7122
4.2058 | -34.0359 | -7.7238
4.2119 | -34.0362 | -7.7335
4.218 | -34.0365 | -7.7437
4.2241 | -34.0368 | -7.7532
4.2303 | -34.0373 | -7.7645
4.2364 | -34.0376 | -7.776
4.2425 | -34.0379 | -7.7865
4.2486 | -34.0381 | -7.7977
4.2547 | -34.0382 | -7.8069
4.2608 | -34.0385 | -7.8169
4.2669 | -34.0388 | -7.8272
4.273 | -34.0391 | -7.838
4.2791 | -34.0393 | -7.8488
4.2852 | -34.0395 | -7.8591
4.2913 | -34.0396 | -7.8683
4.2974 | -34.0399 | -7.8793
4.3035 | -34.0401 | -7.8902
4.3096 | -34.0403 | -7.9006
4.3157 | -34.0405 | -7.9106
4.3218 | -34.0407 | -7.9206
4.3279 | -34.041 | -7.931
4.334 | -34.0412 | -7.9408
4.3401 | -34.0414 | -7.9508
4.3462 | -34.0417 | -7.9616
4.3523 | -34.042 | -7.9713
4.3584 | -34.0422 | -7.9798
4.3645 | -34.0423 | -7.9894
4.3707 | -34.0426 | -7.9987
4.3768 | -34.0428 | -8.0091
4.3829 | -34.0429 | -8.0194
4.389 | -34.0432 | -8.0292
4.3951 | -34.0434 | -8.0387
4.4012 | -34.0436 | -8.0485
4.4073 | -34.0439 | -8.0587
4.4134 | -34.0441 | -8.0688
4.4195 | -34.0443 | -8.0785
4.4256 | -34.0446 | -8.0881
4.4317 | -34.0448 | -8.0976
4.4378 | -34.045 | -8.1071
4.4439 | -34.0452 | -8.1166
4.45 | -34.0453 | -8.1265
4.4561 | -34.0454 | -8.1364
4.4622 | -34.0455 | -8.146
4.4683 | -34.0457 | -8.1552
4.4744 | -34.0459 | -8.1654
4.4805 | -34.0462 | -8.1745
4.4866 | -34.0464 | -8.1838
4.4927 | -34.0466 | -8.194
4.4988 | -34.0467 | -8.2034
4.5049 | -34.047 | -8.2126
4.511 | -34.0472 | -8.2215
4.5172 | -34.0475 | -8.2314
4.5233 | -34.0477 | -8.242
4.5294 | -34.0479 | -8.2514
4.5355 | -34.0482 | -8.2608
4.5416 | -34.0484 | -8.2699
4.5477 | -34.0486 | -8.2792
4.5538 | -34.0488 | -8.2879
4.5599 | -34.0491 | -8.2974
4.566 | -34.0493 | -8.3071
4.5721 | -34.0495 | -8.3163
4.5782 | -34.0498 | -8.3262
4.5843 | -34.05 | -8.3358
4.5904 | -34.0503 | -8.3451
4.5965 | -34.0506 | -8.3539
4.6026 | -34.0509 | -8.3646
4.6087 | -34.0511 | -8.3736
4.6148 | -34.0513 | -8.3843
4.6209 | -34.0517 | -8.3927
4.627 | -34.052 | -8.4028
4.6331 | -34.052 | -8.4117
4.6392 | -34.0523 | -8.4209
4.6453 | -34.0525 | -8.4303
4.6514 | -34.0528 | -8.4399
4.6576 | -34.0529 | -8.4504
4.6637 | -34.0531 | -8.4581
4.6698 | -34.0533 | -8.4666
4.6759 | -34.0534 | -8.4765
4.682 | -34.0536 | -8.4856
4.6881 | -34.0538 | -8.4944
4.6942 | -34.054 | -8.5037
4.7003 | -34.0542 | -8.5131
4.7064 | -34.0545 | -8.5224
4.7125 | -34.0548 | -8.5319
4.7186 | -34.0549 | -8.5403
4.7247 | -34.0553 | -8.55
4.7308 | -34.0556 | -8.5593
4.7369 | -34.0558 | -8.5689
4.743 | -34.056 | -8.5793
4.7491 | -34.0562 | -8.589
4.7552 | -34.0563 | -8.5982
4.7613 | -34.0565 | -8.6079
4.7674 | -34.0568 | -8.6164
4.7735 | -34.0569 | -8.6266
4.7796 | -34.0571 | -8.6364
4.7857 | -34.0572 | -8.6462
4.7918 | -34.0574 | -8.6548
4.7979 | -34.0578 | -8.6638
4.8041 | -34.0579 | -8.6732
4.8102 | -34.0582 | -8.6822
4.8163 | -34.0584 | -8.6917
4.8224 | -34.0585 | -8.7004
4.8285 | -34.0589 | -8.7094
4.8346 | -34.0591 | -8.719
4.8407 | -34.0593 | -8.7283
4.8468 | -34.0596 | -8.7381
4.8529 | -34.0598 | -8.748
4.859 | -34.06 | -8.7575
4.8651 | -34.0602 | -8.7669
4.8712 | -34.0605 | -8.7763
4.8773 | -34.0606 | -8.7859
4.8834 | -34.0609 | -8.7956
4.8895 | -34.0612 | -8.8057
4.8956 | -34.0613 | -8.8157
4.9017 | -34.0615 | -8.8252
4.9078 | -34.0617 | -8.8351
4.9139 | -34.062 | -8.8437
4.92 | -34.0622 | -8.8527
4.9261 | -34.0625 | -8.8616
4.9322 | -34.0629 | -8.8708
4.9383 | -34.0632 | -8.8798
4.9445 | -34.0635 | -8.89
4.9506 | -34.0637 | -8.8989
4.9567 | -34.064 | -8.9082
4.9628 | -34.0642 | -8.9176
4.9689 | -34.0644 | -8.9275
4.975 | -34.0646 | -8.9369
4.9811 | -34.0649 | -8.9467
4.9872 | -34.0652 | -8.9565
4.9933 | -34.0655 | -8.9663
4.9994 | -34.0659 | -8.9762
5.0055 | -34.0662 | -8.9859
5.0116 | -34.0665 | -8.9962
5.0177 | -34.0669 | -9.0063
5.0238 | -34.0672 | -9.0169
5.0299 | -34.0677 | -9.028
5.036 | -34.0679 | -9.0376
5.0421 | -34.0681 | -9.0473
5.0482 | -34.0684 | -9.0576
5.0543 | -34.0688 | -9.0669
5.0604 | -34.0691 | -9.0773
5.0665 | -34.0695 | -9.087
5.0726 | -34.0698 | -9.0975
5.0787 | -34.0701 | -9.1078
5.0848 | -34.0704 | -9.1179
5.091 | -34.0707 | -9.1278
5.0971 | -34.071 | -9.1379
5.1032 | -34.0714 | -9.1484
5.1093 | -34.0717 | -9.1584
5.1154 | -34.0719 | -9.1684
5.1215 | -34.0722 | -9.1785
5.1276 | -34.0725 | -9.1883
5.1337 | -34.073 | -9.1981
5.1398 | -34.0732 | -9.2079
5.1459 | -34.0736 | -9.2184
5.152 | -34.0739 | -9.2295
5.1581 | -34.0741 | -9.2403
5.1642 | -34.0744 | -9.2504
5.1703 | -34.0748 | -9.2604
5.1764 | -34.0751 | -9.2704
5.1825 | -34.0754 | -9.2808
5.1886 | -34.0757 | -9.2906
5.1947 | -34.0759 | -9.3015
5.2008 | -34.0764 | -9.3114
5.2069 | -34.0767 | -9.3226
5.213 | -34.0771 | -9.3323
5.2191 | -34.0775 | -9.3427
5.2252 | -34.0777 | -9.353
5.2314 | -34.0779 | -9.3638
5.2375 | -34.0784 | -9.3738
5.2436 | -34.0786 | -9.3841
5.2497 | -34.079 | -9.394
5.2558 | -34.0794 | -9.4051
5.2619 | -34.0796 | -9.416
5.268 | -34.08 | -9.4266
5.2741 | -34.0804 | -9.4379
5.2802 | -34.0807 | -9.4488
5.2863 | -34.0811 | -9.459
5.2924 | -34.0814 | -9.4703
5.2985 | -34.0818 | -9.4817
5.3046 | -34.0823 | -9.4931
5.3107 | -34.0826 | -9.504
5.3168 | -34.0831 | -9.5145
5.3229 | -34.0834 | -9.5244
5.329 | -34.0838 | -9.5345
5.3351 | -34.0842 | -9.5447
5.3412 | -34.0846 | -9.5549
5.3473 | -34.085 | -9.5659
5.3534 | -34.0854 | -9.5769
5.3595 | -34.0857 | -9.5877
5.3656 | -34.0861 | -9.5987
5.3717 | -34.0864 | -9.6094
5.3779 | -34.0867 | -9.6204
5.384 | -34.087 | -9.6312
5.3901 | -34.0874 | -9.642
5.3962 | -34.0877 | -9.6529
5.4023 | -34.0881 | -9.6637
5.4084 | -34.0885 | -9.6742
5.4145 | -34.0888 | -9.6853
5.4206 | -34.0891 | -9.6967
5.4267 | -34.0895 | -9.7082
5.4328 | -34.0899 | -9.7184
5.4389 | -34.0902 | -9.7287
5.445 | -34.0905 | -9.7401
5.4511 | -34.0909 | -9.7506
5.4572 | -34.0913 | -9.7619
5.4633 | -34.0916 | -9.7723
5.4694 | -34.0919 | -9.7829
5.4755 | -34.0923 | -9.7939
5.4816 | -34.0926 | -9.8046
5.4877 | -34.093 | -9.8154
5.4938 | -34.0933 | -9.8259
5.4999 | -34.0937 | -9.8366
5.506 | -34.0941 | -9.8474
5.5121 | -34.0944 | -9.8586
5.5183 | -34.0947 | -9.8688
5.5244 | -34.0949 | -9.8797
5.5305 | -34.0953 | -9.8906
5.5366 | -34.0957 | -9.9011
5.5427 | -34.096 | -9.9125
5.5488 | -34.0963 | -9.9236
5.5549 | -34.0966 | -9.9344
5.561 | -34.0969 | -9.9458
5.5671 | -34.0974 | -9.9563
5.5732 | -34.0976 | -9.9673
5.5793 | -34.098 | -9.9789
5.5854 | -34.0984 | -9.9899
5.5915 | -34.0986 | -10.001
5.5976 | -34.099 | -10.0121
5.6037 | -34.0993 | -10.0229
5.6098 | -34.0996 | -10.0335
5.6159 | -34.1 | -10.0445
5.622 | -34.1003 | -10.0551
5.6281 | -34.1007 | -10.066
5.6342 | -34.101 | -10.077
5.6403 | -34.1013 | -10.0873
5.6464 | -34.1017 | -10.0982
5.6525 | -34.1019 | -10.1088
5.6586 | -34.1024 | -10.1189
5.6648 | -34.1026 | -10.1303
5.6709 | -34.103 | -10.1415
5.677 | -34.1033 | -10.1531
5.6831 | -34.1036 | -10.1634
5.6892 | -34.104 | -10.1743
5.6953 | -34.1043 | -10.185
5.7014 | -34.1046 | -10.1964
5.7075 | -34.1049 | -10.2075
5.7136 | -34.1053 | -10.2181
5.7197 | -34.1056 | -10.2288
5.7258 | -34.1058 | -10.2394
5.7319 | -34.1062 | -10.2503
5.738 | -34.1065 | -10.2611
5.7441 | -34.1069 | -10.2718
5.7502 | -34.1073 | -10.2822
5.7563 | -34.1076 | -10.2928
5.7624 | -34.1079 | -10.3043
5.7685 | -34.1081 | -10.315
5.7746 | -34.1084 | -10.3256
5.7807 | -34.1087 | -10.336
5.7868 | -34.109 | -10.3468
5.7929 | -34.1093 | -10.3574
5.799 | -34.1096 | -10.3681
5.8052 | -34.1099 | -10.3789
5.8113 | -34.1101 | -10.3897
5.8174 | -34.1104 | -10.4006
5.8235 | -34.1107 | -10.4111
5.8296 | -34.111 | -10.4225
5.8357 | -34.1112 | -10.4332
5.8418 | -34.1116 | -10.4442
5.8479 | -34.1118 | -10.4552
5.854 | -34.1121 | -10.4663
5.8601 | -34.1124 | -10.4768
5.8662 | -34.1127 | -10.4879
5.8723 | -34.1129 | -10.498
5.8784 | -34.1133 | -10.5087
5.8845 | -34.1136 | -10.5196
5.8906 | -34.1138 | -10.5303
5.8967 | -34.1142 | -10.5406
5.9028 | -34.1145 | -10.5513
5.9089 | -34.1148 | -10.5614
5.915 | -34.1151 | -10.5721
5.9211 | -34.1154 | -10.5828
5.9272 | -34.1157 | -10.5934
5.9333 | -34.116 | -10.6034
5.9394 | -34.1163 | -10.6138
5.9455 | -34.1166 | -10.6241
5.9517 | -34.117 | -10.6347
5.9578 | -34.1173 | -10.6457
5.9639 | -34.1177 | -10.656
5.97 | -34.118 | -10.6667
5.9761 | -34.1183 | -10.677
5.9822 | -34.1187 | -10.6878
5.9883 | -34.119 | -10.6982
5.9944 | -34.1193 | -10.7088
6.0005 | -34.1197 | -10.7197
6.0066 | -34.12 | -10.7299
6.0127 | -34.1203 | -10.7398
6.0188 | -34.1206 | -10.7508
6.0249 | -34.1209 | -10.7613
6.031 | -34.1212 | -10.7716
6.0371 | -34.1216 | -10.7825
6.0432 | -34.1219 | -10.793
6.0493 | -34.1223 | -10.8033
6.0554 | -34.1227 | -10.8138
6.0615 | -34.123 | -10.8249
6.0676 | -34.1233 | -10.836
6.0737 | -34.1237 | -10.8468
6.0798 | -34.1239 | -10.8572
6.0859 | -34.1243 | -10.8677
6.0921 | -34.1246 | -10.878
6.0982 | -34.125 | -10.8888
6.1043 | -34.1254 | -10.8991
6.1104 | -34.1257 | -10.9102
6.1165 | -34.1261 | -10.9199
6.1226 | -34.1264 | -10.9299
6.1287 | -34.1268 | -10.9405
6.1348 | -34.1272 | -10.951
6.1409 | -34.1275 | -10.9614
6.147 | -34.1279 | -10.9717
6.1531 | -34.1282 | -10.9822
6.1592 | -34.1285 | -10.9927
6.1653 | -34.1288 | -11.0029
6.1714 | -34.1292 | -11.0134
6.1775 | -34.1296 | -11.0242
6.1836 | -34.1298 | -11.0345
6.1897 | -34.1302 | -11.045
6.1958 | -34.1304 | -11.0557
6.2019 | -34.1309 | -11.0664
6.208 | -34.1312 | -11.0773
6.2141 | -34.1316 | -11.088
6.2202 | -34.1319 | -11.0989
6.2263 | -34.1323 | -11.1093
6.2325 | -34.1327 | -11.1192
6.2386 | -34.1331 | -11.1293
6.2447 | -34.1334 | -11.1401
6.2508 | -34.1337 | -11.1504
6.2569 | -34.1341 | -11.1612
6.263 | -34.1344 | -11.1713
6.2691 | -34.1348 | -11.1817
6.2752 | -34.1352 | -11.192
6.2813 | -34.1355 | -11.2025
6.2874 | -34.1358 | -11.2129
6.2935 | -34.1362 | -11.2231
6.2996 | -34.1365 | -11.2336
6.3057 | -34.1369 | -11.244
6.3118 | -34.1373 | -11.2544
6.3179 | -34.1376 | -11.2648
6.324 | -34.138 | -11.2754
6.3301 | -34.1384 | -11.2857
6.3362 | -34.1388 | -11.2961
6.3423 | -34.1392 | -11.3066
6.3484 | -34.1396 | -11.3173
6.3545 | -34.14 | -11.3271
6.3606 | -34.1404 | -11.3371
6.3667 | -34.1408 | -11.3473
6.3728 | -34.1411 | -11.3577
6.379 | -34.1415 | -11.3681
6.3851 | -34.1419 | -11.3785
6.3912 | -34.1422 | -11.3886
6.3973 | -34.1426 | -11.3988
6.4034 | -34.143 | -11.4091
6.4095 | -34.1433 | -11.4197
6.4156 | -34.1437 | -11.4297
6.4217 | -34.144 | -11.4401
6.4278 | -34.1444 | -11.4503
6.4339 | -34.1448 | -11.4606
6.44 | -34.1452 | -11.4706
6.4461 | -34.1455 | -11.4815
6.4522 | -34.1459 | -11.4912
6.4583 | -34.1462 | -11.5016
6.4644 | -34.1465 | -11.5123
6.4705 | -34.147 | -11.5228
6.4766 | -34.1473 | -11.5336
6.4827 | -34.1477 | -11.5444
6.4888 | -34.148 | -11.5546
6.4949 | -34.1484 | -11.5644
6.501 | -34.1487 | -11.5751
6.5071 | -34.1491 | -11.5856
6.5132 | -34.1496 | -11.5958
6.5194 | -34.1499 | -11.606
6.5255 | -34.1503 | -11.6159
6.5316 | -34.1506 | -11.6264
6.5377 | -34.151 | -11.637
6.5438 | -34.1514 | -11.6472
6.5499 | -34.1518 | -11.6574
6.556 | -34.1522 | -11.6675
6.5621 | -34.1525 | -11.678
6.5682 | -34.1528 | -11.6882
6.5743 | -34.1532 | -11.6993
6.5804 | -34.1535 | -11.7096
6.5865 | -34.1538 | -11.7202
6.5926 | -34.1542 | -11.7305
6.5987 | -34.1546 | -11.7405
6.6048 | -34.1549 | -11.7515
6.6109 | -34.1553 | -11.7617
6.617 | -34.1557 | -11.773
6.6231 | -34.1559 | -11.7835
6.6292 | -34.1564 | -11.7938
6.6353 | -34.1567 | -11.8051
6.6414 | -34.1572 | -11.8158
6.6475 | -34.1576 | -11.8263
6.6536 | -34.1579 | -11.8363
6.6597 | -34.1584 | -11.8464
6.6659 | -34.1587 | -11.8566
6.672 | -34.1591 | -11.867
6.6781 | -34.1595 | -11.8771
6.6842 | -34.1599 | -11.8877
6.6903 | -34.1603 | -11.8979
6.6964 | -34.1606 | -11.9084
6.7025 | -34.161 | -11.9187
6.7086 | -34.1613 | -11.9288
6.7147 | -34.1617 | -11.9394
6.7208 | -34.162 | -11.9502
6.7269 | -34.1624 | -11.9602
6.733 | -34.1627 | -11.9697
6.7391 | -34.1631 | -11.9803
6.7452 | -34.1636 | -11.9901
6.7513 | -34.1639 | -12
6.7574 | -34.1643 | -12.0106
6.7635 | -34.1647 | -12.0204
6.7696 | -34.1652 | -12.0308
6.7757 | -34.1656 | -12.0415
6.7818 | -34.166 | -12.0519
6.7879 | -34.1664 | -12.0623
6.794 | -34.1667 | -12.0724
6.8001 | -34.1671 | -12.0825
6.8063 | -34.1674 | -12.093
6.8124 | -34.1678 | -12.1034
6.8185 | -34.1681 | -12.1137
6.8246 | -34.1685 | -12.1241
6.8307 | -34.1689 | -12.1342
6.8368 | -34.1693 | -12.1449
6.8429 | -34.1697 | -12.1551
6.849 | -34.17 | -12.1654
6.8551 | -34.1704 | -12.1757
6.8612 | -34.1708 | -12.1859
6.8673 | -34.1712 | -12.1959
6.8734 | -34.1716 | -12.207
6.8795 | -34.172 | -12.2176
6.8856 | -34.1723 | -12.2273
6.8917 | -34.1727 | -12.2371
6.8978 | -34.1731 | -12.2473
6.9039 | -34.1734 | -12.2573
6.91 | -34.1738 | -12.2677
6.9161 | -34.1742 | -12.2778
6.9222 | -34.1745 | -12.2881
6.9283 | -34.1749 | -12.298
6.9344 | -34.1753 | -12.3081
6.9405 | -34.1756 | -12.3181
6.9466 | -34.176 | -12.3286
6.9528 | -34.1764 | -12.3389
6.9589 | -34.1767 | -12.349
6.965 | -34.1771 | -12.3593
6.9711 | -34.1775 | -12.3695
6.9772 | -34.178 | -12.3795
6.9833 | -34.1782 | -12.39
6.9894 | -34.1786 | -12.4007
6.9955 | -34.179 | -12.4108
7.0016 | -34.1793 | -12.4215
7.0077 | -34.1797 | -12.432
7.0138 | -34.1801 | -12.4426
7.0199 | -34.1804 | -12.4526
7.026 | -34.1807 | -12.4627
7.0321 | -34.1811 | -12.4729
7.0382 | -34.1814 | -12.4829
7.0443 | -34.1819 | -12.4934
7.0504 | -34.1821 | -12.504
7.0565 | -34.1825 | -12.5144
7.0626 | -34.1829 | -12.5245
7.0687 | -34.1832 | -12.5343
7.0748 | -34.1836 | -12.5443
7.0809 | -34.1839 | -12.5546
7.087 | -34.1843 | -12.5649
7.0932 | -34.1846 | -12.5751
7.0993 | -34.185 | -12.5853
7.1054 | -34.1853 | -12.5953
7.1115 | -34.1856 | -12.6054
7.1176 | -34.1859 | -12.6154
7.1237 | -34.1862 | -12.6256
7.1298 | -34.1865 | -12.6353
7.1359 | -34.1868 | -12.6452
7.142 | -34.1871 | -12.6554
7.1481 | -34.1875 | -12.6652
7.1542 | -34.1878 | -12.6753
7.1603 | -34.1881 | -12.6853
7.1664 | -34.1885 | -12.6955
7.1725 | -34.1889 | -12.7054
7.1786 | -34.1893 | -12.7157
7.1847 | -34.1896 | -12.7257
7.1908 | -34.1899 | -12.7359
7.1969 | -34.1902 | -12.746
7.203 | -34.1906 | -12.7561
7.2091 | -34.1909 | -12.766
7.2152 | -34.1913 | -12.7761
7.2213 | -34.1916 | -12.786
7.2274 | -34.1919 | -12.796
7.2335 | -34.1923 | -12.8061
7.2397 | -34.1926 | -12.8162
7.2458 | -34.1929 | -12.8258
7.2519 | -34.1933 | -12.8358
7.258 | -34.1936 | -12.8461
7.2641 | -34.1939 | -12.8561
7.2702 | -34.1943 | -12.8664
7.2763 | -34.1947 | -12.8763
7.2824 | -34.1951 | -12.8863
7.2885 | -34.1954 | -12.8962
7.2946 | -34.1958 | -12.9065
7.3007 | -34.196 | -12.9165
7.3068 | -34.1963 | -12.9268
7.3129 | -34.1967 | -12.9367
7.319 | -34.197 | -12.9468
7.3251 | -34.1974 | -12.9565
7.3312 | -34.1977 | -12.9662
7.3373 | -34.1981 | -12.9759
7.3434 | -34.1985 | -12.9856
7.3495 | -34.1989 | -12.9952
7.3556 | -34.1992 | -13.0056
7.3617 | -34.1995 | -13.0161
7.3678 | -34.1999 | -13.0262
7.3739 | -34.2002 | -13.0357
7.3801 | -34.2005 | -13.0457
7.3862 | -34.2008 | -13.0556
7.3923 | -34.2012 | -13.0655
7.3984 | -34.2015 | -13.0754
7.4045 | -34.2019 | -13.0852
7.4106 | -34.2022 | -13.095
7.4167 | -34.2025 | -13.1052
7.4228 | -34.2029 | -13.1151
7.4289 | -34.2032 | -13.1248
7.435 | -34.2036 | -13.1344
7.4411 | -34.204 | -13.1441
7.4472 | -34.2043 | -13.1541
7.4533 | -34.2047 | -13.164
7.4594 | -34.205 | -13.1743
7.4655 | -34.2054 | -13.1841
7.4716 | -34.2057 | -13.1943
7.4777 | -34.206 | -13.2037
7.4838 | -34.2064 | -13.2138
7.4899 | -34.2068 | -13.2237
7.496 | -34.2071 | -13.2337
7.5021 | -34.2075 | -13.2435
7.5082 | -34.2078 | -13.2534
7.5143 | -34.2081 | -13.2632
7.5204 | -34.2084 | -13.2729
7.5266 | -34.2087 | -13.2828
7.5327 | -34.2091 | -13.2926
7.5388 | -34.2094 | -13.3024
7.5449 | -34.2097 | -13.3125
7.551 | -34.21 | -13.3226
7.5571 | -34.2104 | -13.3324
7.5632 | -34.2107 | -13.3422
7.5693 | -34.2111 | -13.3523
7.5754 | -34.2115 | -13.3622
7.5815 | -34.2118 | -13.3717
7.5876 | -34.2121 | -13.3815
7.5937 | -34.2124 | -13.3913
7.5998 | -34.2127 | -13.4012
7.6059 | -34.213 | -13.4111
7.612 | -34.2134 | -13.4209
7.6181 | -34.2137 | -13.4307
7.6242 | -34.2141 | -13.4406
7.6303 | -34.2145 | -13.4503
7.6364 | -34.2148 | -13.4604
7.6425 | -34.2151 | -13.4706
7.6486 | -34.2154 | -13.4805
7.6547 | -34.2158 | -13.4905
7.6608 | -34.2162 | -13.5003
7.667 | -34.2165 | -13.5096
7.6731 | -34.2168 | -13.5197
7.6792 | -34.2171 | -13.5297
7.6853 | -34.2174 | -13.5394
7.6914 | -34.2176 | -13.5495
7.6975 | -34.218 | -13.5595
7.7036 | -34.2183 | -13.5695
7.7097 | -34.2186 | -13.5793
7.7158 | -34.219 | -13.589
7.7219 | -34.2192 | -13.5988
7.728 | -34.2195 | -13.6088
7.7341 | -34.2198 | -13.6188
7.7402 | -34.2202 | -13.6285
7.7463 | -34.2205 | -13.6384
7.7524 | -34.2208 | -13.6486
7.7585 | -34.2211 | -13.6586
7.7646 | -34.2214 | -13.6682
7.7707 | -34.2217 | -13.6779
7.7768 | -34.2219 | -13.6876
7.7829 | -34.2222 | -13.6976
7.789 | -34.2225 | -13.7072
7.7951 | -34.2228 | -13.7174
7.8012 | -34.2232 | -13.7277
7.8073 | -34.2235 | -13.7374
7.8135 | -34.2237 | -13.7474
7.8196 | -34.2241 | -13.7572
7.8257 | -34.2244 | -13.7675
7.8318 | -34.2247 | -13.7778
7.8379 | -34.2251 | -13.7872
7.844 | -34.2253 | -13.7966
7.8501 | -34.2255 | -13.8068
7.8562 | -34.2258 | -13.8173
7.8623 | -34.2262 | -13.8272
7.8684 | -34.2264 | -13.8368
7.8745 | -34.2268 | -13.8466
7.8806 | -34.227 | -13.856
7.8867 | -34.2274 | -13.8657
7.8928 | -34.2277 | -13.8754
7.8989 | -34.2281 | -13.8848
7.905 | -34.2284 | -13.8951
7.9111 | -34.2287 | -13.905
7.9172 | -34.229 | -13.9146
7.9233 | -34.2292 | -13.9245
7.9294 | -34.2296 | -13.9345
7.9355 | -34.2299 | -13.9439
7.9416 | -34.2301 | -13.9536
7.9477 | -34.2305 | -13.963
7.9539 | -34.2308 | -13.9725
7.96 | -34.2311 | -13.9826
7.9661 | -34.2314 | -13.9925
7.9722 | -34.2318 | -14.0024
7.9783 | -34.232 | -14.0122
7.9844 | -34.2324 | -14.0225
7.9905 | -34.2327 | -14.0322
7.9966 | -34.2329 | -14.042
8.0027 | -34.2332 | -14.0522
8.0088 | -34.2335 | -14.0616
8.0149 | -34.2338 | -14.0712
8.021 | -34.2341 | -14.081
8.0271 | -34.2344 | -14.0913
8.0332 | -34.2347 | -14.1011
8.0393 | -34.235 | -14.1112
8.0454 | -34.2353 | -14.1212
8.0515 | -34.2356 | -14.131
8.0576 | -34.2359 | -14.1411
8.0637 | -34.2362 | -14.1504
8.0698 | -34.2365 | -14.1598
8.0759 | -34.2368 | -14.1693
8.082 | -34.2371 | -14.1793
8.0881 | -34.2374 | -14.189
8.0942 | -34.2377 | -14.1993
8.1004 | -34.238 | -14.2095
8.1065 | -34.2383 | -14.2193
8.1126 | -34.2386 | -14.229
8.1187 | -34.239 | -14.239
8.1248 | -34.2392 | -14.249
8.1309 | -34.2395 | -14.2588
8.137 | -34.2398 | -14.2685
8.1431 | -34.2402 | -14.2785
8.1492 | -34.2405 | -14.2885
8.1553 | -34.2408 | -14.2983
8.1614 | -34.2411 | -14.3084
8.1675 | -34.2413 | -14.3181
8.1736 | -34.2416 | -14.3282
8.1797 | -34.2419 | -14.3383
8.1858 | -34.2422 | -14.348
8.1919 | -34.2424 | -14.3578
8.198 | -34.2427 | -14.3673
8.2041 | -34.2431 | -14.3775
8.2102 | -34.2434 | -14.3874
8.2163 | -34.2437 | -14.3971
8.2224 | -34.244 | -14.4072
8.2285 | -34.2443 | -14.4168
8.2346 | -34.2447 | -14.4266
8.2408 | -34.245 | -14.4365
8.2469 | -34.2453 | -14.4463
8.253 | -34.2455 | -14.456
8.2591 | -34.2458 | -14.4656
8.2652 | -34.2462 | -14.4753
8.2713 | -34.2466 | -14.4852
8.2774 | -34.2469 | -14.4951
8.2835 | -34.2472 | -14.5054
8.2896 | -34.2475 | -14.5147
8.2957 | -34.2478 | -14.5245
8.3018 | -34.2481 | -14.5343
8.3079 | -34.2484 | -14.5441
8.314 | -34.2487 | -14.5538
8.3201 | -34.2491 | -14.5636
8.3262 | -34.2493 | -14.5733
8.3323 | -34.2497 | -14.5831
8.3384 | -34.25 | -14.5934
8.3445 | -34.2503 | -14.6027
8.3506 | -34.2506 | -14.6128
8.3567 | -34.2509 | -14.6229
8.3628 | -34.2512 | -14.6326
8.3689 | -34.2515 | -14.6425
8.375 | -34.2518 | -14.6523
8.3812 | -34.2521 | -14.6621
8.3873 | -34.2524 | -14.672
8.3934 | -34.2527 | -14.6818
8.3995 | -34.253 | -14.6917
8.4056 | -34.2533 | -14.7017
8.4117 | -34.2536 | -14.7118
8.4178 | -34.254 | -14.7212
8.4239 | -34.2543 | -14.7308
8.43 | -34.2546 | -14.7406
8.4361 | -34.2549 | -14.7508
8.4422 | -34.2552 | -14.7606
8.4483 | -34.2555 | -14.7703
8.4544 | -34.2558 | -14.7805
8.4605 | -34.2561 | -14.7902
8.4666 | -34.2564 | -14.8002
8.4727 | -34.2566 | -14.8099
8.4788 | -34.257 | -14.8196
8.4849 | -34.2573 | -14.8297
8.491 | -34.2575 | -14.8394
8.4971 | -34.2579 | -14.8493
8.5032 | -34.2582 | -14.8591
8.5093 | -34.2585 | -14.8689
8.5154 | -34.2588 | -14.8785
8.5215 | -34.2591 | -14.8884
8.5277 | -34.2594 | -14.8983
8.5338 | -34.2597 | -14.9079
8.5399 | -34.26 | -14.918
8.546 | -34.2603 | -14.9275
8.5521 | -34.2606 | -14.9375
8.5582 | -34.2609 | -14.9474
8.5643 | -34.2612 | -14.957
8.5704 | -34.2615 | -14.9665
8.5765 | -34.2618 | -14.9763
8.5826 | -34.2622 | -14.9861
8.5887 | -34.2625 | -14.9957
8.5948 | -34.2628 | -15.0052
8.6009 | -34.2631 | -15.0149
8.607 | -34.2634 | -15.0249
8.6131 | -34.2637 | -15.035
8.6192 | -34.264 | -15.0444
8.6253 | -34.2644 | -15.0544
8.6314 | -34.2647 | -15.0641
8.6375 | -34.265 | -15.0741
8.6436 | -34.2654 | -15.084
8.6497 | -34.2657 | -15.0938
8.6558 | -34.266 | -15.1034
8.6619 | -34.2662 | -15.1134
8.6681 | -34.2665 | -15.1231
8.6742 | -34.2669 | -15.1331
8.6803 | -34.2672 | -15.1432
8.6864 | -34.2676 | -15.1532
8.6925 | -34.2679 | -15.163
8.6986 | -34.2682 | -15.1732
8.7047 | -34.2685 | -15.1831
8.7108 | -34.2688 | -15.193
8.7169 | -34.2691 | -15.2026
8.723 | -34.2695 | -15.2124
8.7291 | -34.2698 | -15.2221
8.7352 | -34.2702 | -15.2322
8.7413 | -34.2704 | -15.2419
8.7474 | -34.2708 | -15.2514
8.7535 | -34.2711 | -15.2613
8.7596 | -34.2715 | -15.2717
8.7657 | -34.2718 | -15.2815
8.7718 | -34.2721 | -15.2911
8.7779 | -34.2724 | -15.3015
8.784 | -34.2727 | -15.3111
8.7901 | -34.2729 | -15.3214
8.7962 | -34.2733 | -15.3313
8.8023 | -34.2737 | -15.3408
8.8084 | -34.274 | -15.3508
8.8146 | -34.2743 | -15.3605
8.8207 | -34.2747 | -15.3704
8.8268 | -34.275 | -15.3801
8.8329 | -34.2753 | -15.39
8.839 | -34.2757 | -15.3999
8.8451 | -34.276 | -15.4096
8.8512 | -34.2764 | -15.4194
8.8573 | -34.2767 | -15.4291
8.8634 | -34.2771 | -15.4388
8.8695 | -34.2774 | -15.4487
8.8756 | -34.2777 | -15.4583
8.8817 | -34.2781 | -15.4684
8.8878 | -34.2783 | -15.4785
8.8939 | -34.2786 | -15.488
8.9 | -34.279 | -15.4975
8.9061 | -34.2793 | -15.5074
8.9122 | -34.2797 | -15.5171
8.9183 | -34.28 | -15.5268
8.9244 | -34.2803 | -15.5364
8.9305 | -34.2807 | -15.5462
8.9366 | -34.281 | -15.5559
8.9427 | -34.2814 | -15.5658
8.9488 | -34.2817 | -15.5755
8.955 | -34.2821 | -15.5853
8.9611 | -34.2824 | -15.5952
8.9672 | -34.2828 | -15.6054
8.9733 | -34.2831 | -15.6151
8.9794 | -34.2835 | -15.6251
8.9855 | -34.2838 | -15.6348
8.9916 | -34.2841 | -15.6445
8.9977 | -34.2845 | -15.6542
9.0038 | -34.2849 | -15.6642
9.0099 | -34.2852 | -15.6741
9.016 | -34.2856 | -15.6839
9.0221 | -34.2859 | -15.6939
9.0282 | -34.2863 | -15.7038
9.0343 | -34.2867 | -15.7139
9.0404 | -34.2871 | -15.7236
9.0465 | -34.2874 | -15.7338
9.0526 | -34.2878 | -15.7435
9.0587 | -34.2882 | -15.7533
9.0648 | -34.2886 | -15.763
9.0709 | -34.2889 | -15.7726
9.077 | -34.2892 | -15.783
9.0831 | -34.2896 | -15.7928
9.0892 | -34.2899 | -15.8026
9.0953 | -34.2902 | -15.8126
9.1015 | -34.2906 | -15.8224
9.1076 | -34.291 | -15.8324
9.1137 | -34.2913 | -15.8421
9.1198 | -34.2917 | -15.8516
9.1259 | -34.2921 | -15.8615
9.132 | -34.2925 | -15.8716
9.1381 | -34.2929 | -15.8813
9.1442 | -34.2933 | -15.8907
9.1503 | -34.2936 | -15.9007
9.1564 | -34.294 | -15.9105
9.1625 | -34.2943 | -15.9204
9.1686 | -34.2947 | -15.9302
9.1747 | -34.2951 | -15.9401
9.1808 | -34.2955 | -15.9502
9.1869 | -34.2958 | -15.9603
9.193 | -34.2962 | -15.9703
9.1991 | -34.2965 | -15.9798
9.2052 | -34.2969 | -15.9895
9.2113 | -34.2973 | -15.9996
9.2174 | -34.2976 | -16.0101
9.2235 | -34.298 | -16.0196
9.2296 | -34.2985 | -16.0296
9.2357 | -34.2988 | -16.0395
9.2419 | -34.2992 | -16.0493
9.248 | -34.2996 | -16.0594
9.2541 | -34.3 | -16.0694
9.2602 | -34.3003 | -16.0794
9.2663 | -34.3007 | -16.089
9.2724 | -34.301 | -16.0989
9.2785 | -34.3014 | -16.1087
9.2846 | -34.3018 | -16.1186
9.2907 | -34.3022 | -16.1284
9.2968 | -34.3025 | -16.1384
9.3029 | -34.3028 | -16.1484
9.309 | -34.3032 | -16.1582
9.3151 | -34.3036 | -16.1679
9.3212 | -34.304 | -16.1779
9.3273 | -34.3045 | -16.1876
9.3334 | -34.3049 | -16.1973
9.3395 | -34.3053 | -16.207
9.3456 | -34.3057 | -16.2166
9.3517 | -34.3061 | -16.226
9.3578 | -34.3065 | -16.2357
9.3639 | -34.3069 | -16.2457
9.37 | -34.3073 | -16.2556
9.3761 | -34.3077 | -16.2653
9.3822 | -34.308 | -16.2751
9.3884 | -34.3085 | -16.2849
9.3945 | -34.3088 | -16.2948
9.4006 | -34.3092 | -16.3047
9.4067 | -34.3096 | -16.3147
9.4128 | -34.31 | -16.3244
9.4189 | -34.3104 | -16.3344
9.425 | -34.3108 | -16.3441
9.4311 | -34.3112 | -16.3539
9.4372 | -34.3116 | -16.3637
9.4433 | -34.312 | -16.3737
9.4494 | -34.3123 | -16.3835
9.4555 | -34.3127 | -16.3933
9.4616 | -34.3132 | -16.4031
9.4677 | -34.3136 | -16.4127
9.4738 | -34.314 | -16.4229
9.4799 | -34.3143 | -16.4327
9.486 | -34.3146 | -16.4428
9.4921 | -34.3151 | -16.4531
9.4982 | -34.3155 | -16.4625
9.5043 | -34.3159 | -16.472
9.5104 | -34.3163 | -16.4817
9.5165 | -34.3167 | -16.4916
9.5226 | -34.3171 | -16.5015
9.5288 | -34.3175 | -16.5113
9.5349 | -34.3179 | -16.5212
9.541 | -34.3183 | -16.5309
9.5471 | -34.3187 | -16.5408
9.5532 | -34.3191 | -16.5507
9.5593 | -34.3195 | -16.5605
9.5654 | -34.3199 | -16.5704
9.5715 | -34.3203 | -16.5804
9.5776 | -34.3207 | -16.5905
9.5837 | -34.3211 | -16.6
9.5898 | -34.3214 | -16.61
9.5959 | -34.3218 | -16.6197
9.602 | -34.3222 | -16.6296
9.6081 | -34.3226 | -16.6396
9.6142 | -34.323 | -16.6493
9.6203 | -34.3234 | -16.6592
9.6264 | -34.3237 | -16.6686
9.6325 | -34.324 | -16.6787
9.6386 | -34.3245 | -16.6884
9.6447 | -34.3248 | -16.6981
9.6508 | -34.3253 | -16.708
9.6569 | -34.3257 | -16.7178
9.663 | -34.3261 | -16.7277
9.6691 | -34.3265 | -16.7376
9.6753 | -34.327 | -16.7474
9.6814 | -34.3274 | -16.757
9.6875 | -34.3278 | -16.7666
9.6936 | -34.3282 | -16.7761
9.6997 | -34.3286 | -16.7862
9.7058 | -34.3291 | -16.796
9.7119 | -34.3294 | -16.8057
9.718 | -34.3298 | -16.8159
9.7241 | -34.3302 | -16.8257
9.7302 | -34.3306 | -16.8355
9.7363 | -34.331 | -16.8451
9.7424 | -34.3314 | -16.8548
9.7485 | -34.3318 | -16.8645
9.7546 | -34.3322 | -16.874
9.7607 | -34.3326 | -16.884
9.7668 | -34.333 | -16.8943
9.7729 | -34.3335 | -16.9038
9.779 | -34.3339 | -16.9136
9.7851 | -34.3343 | -16.9235
9.7912 | -34.3347 | -16.9339
9.7973 | -34.335 | -16.9439
9.8034 | -34.3355 | -16.9538
9.8095 | -34.3359 | -16.9635
9.8157 | -34.3363 | -16.9733
9.8218 | -34.3367 | -16.9828
9.8279 | -34.3371 | -16.9926
9.834 | -34.3375 | -17.0023
9.8401 | -34.3379 | -17.012
9.8462 | -34.3384 | -17.0215
9.8523 | -34.3388 | -17.0312
9.8584 | -34.3392 | -17.0412
9.8645 | -34.3396 | -17.0511
9.8706 | -34.3399 | -17.0608
9.8767 | -34.3403 | -17.0702
9.8828 | -34.3408 | -17.0796
9.8889 | -34.3412 | -17.0896
9.895 | -34.3415 | -17.099
9.9011 | -34.3419 | -17.1083
9.9072 | -34.3424 | -17.1177
9.9133 | -34.3428 | -17.1273
9.9194 | -34.3432 | -17.1368
9.9255 | -34.3436 | -17.1467
9.9316 | -34.3441 | -17.1562
9.9377 | -34.3445 | -17.1658
9.9438 | -34.3449 | -17.1754
9.9499 | -34.3453 | -17.1852
9.956 | -34.3457 | -17.1948
9.9622 | -34.3461 | -17.2044
9.9683 | -34.3465 | -17.2142
9.9744 | -34.3469 | -17.2238
9.9805 | -34.3474 | -17.2336
9.9866 | -34.3477 | -17.2429
9.9927 | -34.3482 | -17.2528
9.9988 | -34.3486 | -17.2624
10.0049 | -34.349 | -17.272
10.011 | -34.3495 | -17.2816
10.0171 | -34.3498 | -17.2913
10.0232 | -34.3503 | -17.3007
10.0293 | -34.3507 | -17.3102
10.0354 | -34.3512 | -17.3196
10.0415 | -34.3516 | -17.3292
10.0476 | -34.352 | -17.3392
10.0537 | -34.3524 | -17.3491
10.0598 | -34.3528 | -17.3589
10.0659 | -34.3532 | -17.3685
10.072 | -34.3536 | -17.3781
10.0781 | -34.3541 | -17.3877
10.0842 | -34.3545 | -17.3975
10.0903 | -34.3549 | -17.4074
10.0964 | -34.3554 | -17.4168
10.1026 | -34.3558 | -17.4264
10.1087 | -34.3562 | -17.4361
10.1148 | -34.3567 | -17.4458
10.1209 | -34.3571 | -17.4553
10.127 | -34.3575 | -17.4647
10.1331 | -34.3579 | -17.4742
10.1392 | -34.3584 | -17.4837
10.1453 | -34.3588 | -17.4934
10.1514 | -34.3592 | -17.5028
10.1575 | -34.3596 | -17.5123
10.1636 | -34.3601 | -17.5212
10.1697 | -34.3604 | -17.5308
10.1758 | -34.3609 | -17.5402
10.1819 | -34.3613 | -17.5499
10.188 | -34.3617 | -17.5592
10.1941 | -34.3621 | -17.5689
10.2002 | -34.3624 | -17.5779
10.2063 | -34.3629 | -17.5877
10.2124 | -34.3633 | -17.5973
10.2185 | -34.3637 | -17.6069
10.2246 | -34.3641 | -17.6164
10.2307 | -34.3645 | -17.6261
10.2368 | -34.365 | -17.6347
10.2429 | -34.3654 | -17.644
10.2491 | -34.3658 | -17.6533
10.2552 | -34.3662 | -17.6633
10.2613 | -34.3667 | -17.6731
10.2674 | -34.367 | -17.6828
10.2735 | -34.3674 | -17.6926
10.2796 | -34.3678 | -17.7023
10.2857 | -34.3682 | -17.7119
10.2918 | -34.3686 | -17.7215
10.2979 | -34.3691 | -17.7312
10.304 | -34.3696 | -17.741
10.3101 | -34.37 | -17.7505
10.3162 | -34.3704 | -17.7601
10.3223 | -34.3708 | -17.7696
10.3284 | -34.3712 | -17.779
10.3345 | -34.3716 | -17.7884
10.3406 | -34.372 | -17.7978
10.3467 | -34.3725 | -17.8073
10.3528 | -34.3729 | -17.8165
10.3589 | -34.3732 | -17.8262
10.365 | -34.3737 | -17.836
10.3711 | -34.374 | -17.8457
10.3772 | -34.3744 | -17.8553
10.3833 | -34.3748 | -17.8648
10.3895 | -34.3752 | -17.8741
10.3956 | -34.3757 | -17.8834
10.4017 | -34.3761 | -17.8929
10.4078 | -34.3765 | -17.9024
10.4139 | -34.3769 | -17.9119
10.42 | -34.3773 | -17.9215
10.4261 | -34.3778 | -17.9309
10.4322 | -34.3782 | -17.9402
10.4383 | -34.3786 | -17.9498
10.4444 | -34.379 | -17.9588
10.4505 | -34.3794 | -17.9683
10.4566 | -34.3798 | -17.9777
10.4627 | -34.3802 | -17.9871
10.4688 | -34.3807 | -17.9961
10.4749 | -34.3811 | -18.0054
10.481 | -34.3815 | -18.0149
10.4871 | -34.3819 | -18.0245
10.4932 | -34.3824 | -18.0338
10.4993 | -34.3828 | -18.0432
10.5054 | -34.3832 | -18.0526
10.5115 | -34.3837 | -18.0621
10.5176 | -34.3841 | -18.0713
10.5237 | -34.3846 | -18.0807
10.5298 | -34.3849 | -18.0899
10.536 | -34.3854 | -18.0995
10.5421 | -34.3858 | -18.1085
10.5482 | -34.3862 | -18.1182
10.5543 | -34.3867 | -18.1276
10.5604 | -34.3871 | -18.1366
10.5665 | -34.3875 | -18.1459
10.5726 | -34.3879 | -18.1551
10.5787 | -34.3883 | -18.1646
10.5848 | -34.3888 | -18.1735
10.5909 | -34.3892 | -18.183
10.597 | -34.3896 | -18.1924
10.6031 | -34.39 | -18.202
10.6092 | -34.3904 | -18.2117
10.6153 | -34.3908 | -18.2213
10.6214 | -34.3913 | -18.2305
10.6275 | -34.3917 | -18.2398
10.6336 | -34.3921 | -18.2495
10.6397 | -34.3925 | -18.2587
10.6458 | -34.3929 | -18.2678
10.6519 | -34.3933 | -18.2771
10.658 | -34.3937 | -18.2862
10.6641 | -34.3941 | -18.2956
10.6702 | -34.3946 | -18.3051
10.6764 | -34.395 | -18.3144
10.6825 | -34.3955 | -18.3239
10.6886 | -34.3959 | -18.3334
10.6947 | -34.3963 | -18.3425
10.7008 | -34.3967 | -18.3519
10.7069 | -34.3972 | -18.3613
10.713 | -34.3976 | -18.3706
10.7191 | -34.398 | -18.3801
10.7252 | -34.3985 | -18.3892
10.7313 | -34.3989 | -18.3984
10.7374 | -34.3993 | -18.4075
10.7435 | -34.3997 | -18.4171
10.7496 | -34.4002 | -18.4262
10.7557 | -34.4006 | -18.4356
10.7618 | -34.401 | -18.4448
10.7679 | -34.4014 | -18.4542
10.774 | -34.4018 | -18.4635
10.7801 | -34.4022 | -18.4732
10.7862 | -34.4027 | -18.4825
10.7923 | -34.4031 | -18.4919
10.7984 | -34.4035 | -18.5012
10.8045 | -34.4039 | -18.5107
10.8106 | -34.4044 | -18.5201
10.8168 | -34.4048 | -18.5294
10.8229 | -34.4052 | -18.5387
10.829 | -34.4056 | -18.5481
10.8351 | -34.406 | -18.5574
10.8412 | -34.4064 | -18.5666
10.8473 | -34.4068 | -18.5756
10.8534 | -34.4072 | -18.5851
10.8595 | -34.4077 | -18.5942
10.8656 | -34.4081 | -18.6033
10.8717 | -34.4086 | -18.6122
10.8778 | -34.409 | -18.6214
10.8839 | -34.4094 | -18.6306
10.89 | -34.4098 | -18.6397
10.8961 | -34.4103 | -18.649
10.9022 | -34.4107 | -18.6582
10.9083 | -34.4111 | -18.6678
10.9144 | -34.4116 | -18.6769
10.9205 | -34.412 | -18.6862
10.9266 | -34.4125 | -18.6958
10.9327 | -34.413 | -18.705
10.9388 | -34.4134 | -18.7143
10.9449 | -34.4138 | -18.7237
10.951 | -34.4142 | -18.733
10.9571 | -34.4146 | -18.7423
10.9633 | -34.415 | -18.752
10.9694 | -34.4155 | -18.7616
10.9755 | -34.4159 | -18.7707
10.9816 | -34.4163 | -18.78
10.9877 | -34.4168 | -18.7893
10.9938 | -34.4171 | -18.7988
10.9999 | -34.4176 | -18.8079
11.006 | -34.4181 | -18.8171
11.0121 | -34.4185 | -18.8267
11.0182 | -34.4189 | -18.8357
11.0243 | -34.4193 | -18.845
11.0304 | -34.4197 | -18.8542
11.0365 | -34.4202 | -18.8633
11.0426 | -34.4206 | -18.8723
11.0487 | -34.421 | -18.8816
11.0548 | -34.4215 | -18.8907
11.0609 | -34.4219 | -18.9002
11.067 | -34.4223 | -18.9097
11.0731 | -34.4227 | -18.919
11.0792 | -34.4232 | -18.9282
11.0853 | -34.4236 | -18.9372
11.0914 | -34.424 | -18.9464
11.0975 | -34.4245 | -18.9559
11.1037 | -34.4249 | -18.9648
11.1098 | -34.4253 | -18.9742
11.1159 | -34.4258 | -18.9837
11.122 | -34.4262 | -18.9931
11.1281 | -34.4266 | -19.0021
11.1342 | -34.427 | -19.0111
11.1403 | -34.4274 | -19.0206
11.1464 | -34.4279 | -19.0298
11.1525 | -34.4283 | -19.0392
11.1586 | -34.4287 | -19.0489
11.1647 | -34.4292 | -19.0579
11.1708 | -34.4296 | -19.0671
11.1769 | -34.43 | -19.0759
11.183 | -34.4305 | -19.0852
11.1891 | -34.431 | -19.0947
11.1952 | -34.4314 | -19.1042
11.2013 | -34.4318 | -19.1134
11.2074 | -34.4322 | -19.1225
11.2135 | -34.4325 | -19.1319
11.2196 | -34.4329 | -19.1409
11.2257 | -34.4333 | -19.1499
11.2318 | -34.4337 | -19.1594
11.2379 | -34.4341 | -19.1685
11.244 | -34.4345 | -19.1776
11.2502 | -34.4349 | -19.1867
11.2563 | -34.4354 | -19.196
11.2624 | -34.4358 | -19.2052
11.2685 | -34.4362 | -19.2142
11.2746 | -34.4366 | -19.2232
11.2807 | -34.437 | -19.2329
11.2868 | -34.4374 | -19.2425
11.2929 | -34.4378 | -19.2514
11.299 | -34.4382 | -19.2609
11.3051 | -34.4386 | -19.2703
11.3112 | -34.439 | -19.2794
11.3173 | -34.4395 | -19.2887
11.3234 | -34.4399 | -19.2981
11.3295 | -34.4403 | -19.307
11.3356 | -34.4408 | -19.3161
11.3417 | -34.4412 | -19.3254
11.3478 | -34.4416 | -19.3345
11.3539 | -34.442 | -19.3436
11.36 | -34.4424 | -19.3527
11.3661 | -34.4428 | -19.3618
11.3722 | -34.4433 | -19.3709
11.3783 | -34.4437 | -19.3801
11.3844 | -34.4441 | -19.3893
11.3906 | -34.4445 | -19.3988
11.3967 | -34.4449 | -19.4079
11.4028 | -34.4453 | -19.4175
11.4089 | -34.4457 | -19.4266
11.415 | -34.4461 | -19.4358
11.4211 | -34.4465 | -19.4448
11.4272 | -34.4469 | -19.4541
11.4333 | -34.4473 | -19.4632
11.4394 | -34.4478 | -19.4727
11.4455 | -34.4482 | -19.4818
11.4516 | -34.4486 | -19.4908
11.4577 | -34.449 | -19.5001
11.4638 | -34.4495 | -19.5088
11.4699 | -34.4498 | -19.5183
11.476 | -34.4503 | -19.5278
11.4821 | -34.4507 | -19.5366
11.4882 | -34.4511 | -19.546
11.4943 | -34.4515 | -19.5549
11.5004 | -34.452 | -19.5636
11.5065 | -34.4525 | -19.5728
11.5126 | -34.4529 | -19.5818
11.5187 | -34.4534 | -19.5909
11.5248 | -34.4538 | -19.6001
11.5309 | -34.4542 | -19.6093
11.5371 | -34.4545 | -19.6185
11.5432 | -34.4549 | -19.6279
11.5493 | -34.4553 | -19.6372
11.5554 | -34.4557 | -19.6462
11.5615 | -34.4561 | -19.6558
11.5676 | -34.4566 | -19.6652
11.5737 | -34.457 | -19.6738
11.5798 | -34.4574 | -19.6828
11.5859 | -34.4579 | -19.6918
11.592 | -34.4583 | -19.701
11.5981 | -34.4587 | -19.71
11.6042 | -34.4591 | -19.7195
11.6103 | -34.4595 | -19.7286
11.6164 | -34.46 | -19.7375
11.6225 | -34.4604 | -19.7468
11.6286 | -34.4608 | -19.7561
11.6347 | -34.4612 | -19.7656
11.6408 | -34.4616 | -19.7748
11.6469 | -34.4621 | -19.7842
11.653 | -34.4625 | -19.793
11.6591 | -34.463 | -19.8023
11.6652 | -34.4634 | -19.8116
11.6713 | -34.4638 | -19.8207
11.6775 | -34.4643 | -19.8299
11.6836 | -34.4647 | -19.8392
11.6897 | -34.4651 | -19.8483
11.6958 | -34.4656 | -19.8579
11.7019 | -34.466 | -19.867
11.708 | -34.4664 | -19.8764
11.7141 | -34.4669 | -19.8856
11.7202 | -34.4673 | -19.8945
11.7263 | -34.4677 | -19.9037
11.7324 | -34.4681 | -19.913
11.7385 | -34.4685 | -19.9224
11.7446 | -34.4689 | -19.9317
11.7507 | -34.4693 | -19.9408
11.7568 | -34.4697 | -19.9501
11.7629 | -34.4702 | -19.9593
11.769 | -34.4706 | -19.9684
11.7751 | -34.4711 | -19.9772
11.7812 | -34.4715 | -19.9863
11.7873 | -34.4719 | -19.9956
11.7934 | -34.4724 | -20.0047
11.7995 | -34.4728 | -20.0134
11.8056 | -34.4732 | -20.023
11.8117 | -34.4736 | -20.0321
11.8178 | -34.474 | -20.0413
11.824 | -34.4745 | -20.0502
11.8301 | -34.475 | -20.059
11.8362 | -34.4753 | -20.0678
11.8423 | -34.4757 | -20.077
11.8484 | -34.4762 | -20.0862
11.8545 | -34.4766 | -20.0956
11.8606 | -34.4771 | -20.1048
11.8667 | -34.4775 | -20.114
11.8728 | -34.478 | -20.1231
11.8789 | -34.4784 | -20.1327
11.885 | -34.4787 | -20.1418
11.8911 | -34.4792 | -20.1512
11.8972 | -34.4796 | -20.1599
11.9033 | -34.48 | -20.1686
11.9094 | -34.4804 | -20.1774
11.9155 | -34.4808 | -20.1866
11.9216 | -34.4813 | -20.1957
11.9277 | -34.4818 | -20.2046
11.9338 | -34.4821 | -20.2137
11.9399 | -34.4826 | -20.2232
11.946 | -34.483 | -20.2323
11.9521 | -34.4835 | -20.2413
11.9582 | -34.4839 | -20.2502
11.9644 | -34.4844 | -20.2591
11.9705 | -34.4848 | -20.2684
11.9766 | -34.4852 | -20.2774
11.9827 | -34.4856 | -20.2865
11.9888 | -34.486 | -20.2954
11.9949 | -34.4864 | -20.3045
12.001 | -34.4869 | -20.3136
12.0071 | -34.4874 | -20.3225
12.0132 | -34.4878 | -20.3313
12.0193 | -34.4882 | -20.3407
12.0254 | -34.4886 | -20.3495
12.0315 | -34.4891 | -20.3585
12.0376 | -34.4895 | -20.3673
12.0437 | -34.4899 | -20.3761
12.0498 | -34.4903 | -20.3854
12.0559 | -34.4908 | -20.3945
12.062 | -34.4912 | -20.4036
12.0681 | -34.4916 | -20.4126
12.0742 | -34.4921 | -20.4218
12.0803 | -34.4925 | -20.4308
12.0864 | -34.493 | -20.44
12.0925 | -34.4934 | -20.4491
12.0986 | -34.4939 | -20.4579
12.1047 | -34.4943 | -20.467
12.1109 | -34.4948 | -20.4762
12.117 | -34.4952 | -20.4852
12.1231 | -34.4957 | -20.4943
12.1292 | -34.4961 | -20.5035
12.1353 | -34.4965 | -20.5124
12.1414 | -34.497 | -20.5217
12.1475 | -34.4974 | -20.5309
12.1536 | -34.4978 | -20.54
12.1597 | -34.4982 | -20.5491
12.1658 | -34.4987 | -20.5581
12.1719 | -34.4992 | -20.5672
12.178 | -34.4996 | -20.5764
12.1841 | -34.5 | -20.5857
12.1902 | -34.5005 | -20.5947
12.1963 | -34.5009 | -20.6037
12.2024 | -34.5014 | -20.6127
12.2085 | -34.5018 | -20.6217
12.2146 | -34.5022 | -20.6306
12.2207 | -34.5027 | -20.6397
12.2268 | -34.5031 | -20.6487
12.2329 | -34.5035 | -20.6578
12.239 | -34.5039 | -20.6665
12.2451 | -34.5044 | -20.6753
12.2513 | -34.5049 | -20.6844
12.2574 | -34.5053 | -20.6938
12.2635 | -34.5057 | -20.7024
12.2696 | -34.5061 | -20.7111
12.2757 | -34.5066 | -20.7198
12.2818 | -34.507 | -20.7287
12.2879 | -34.5074 | -20.7375
12.294 | -34.5079 | -20.7463
12.3001 | -34.5083 | -20.7556
12.3062 | -34.5088 | -20.7647
12.3123 | -34.5092 | -20.7737
12.3184 | -34.5097 | -20.7825
12.3245 | -34.5101 | -20.7916
12.3306 | -34.5106 | -20.8003
12.3367 | -34.511 | -20.8091
12.3428 | -34.5115 | -20.8179
12.3489 | -34.5119 | -20.8264
12.355 | -34.5123 | -20.8353
12.3611 | -34.5128 | -20.8443
12.3672 | -34.5132 | -20.8531
12.3733 | -34.5137 | -20.8624
12.3794 | -34.5141 | -20.8709
12.3855 | -34.5145 | -20.8797
12.3916 | -34.5149 | -20.8888
12.3978 | -34.5154 | -20.898
12.4039 | -34.5157 | -20.907
12.41 | -34.5162 | -20.9161
12.4161 | -34.5167 | -20.9247
12.4222 | -34.5171 | -20.9336
12.4283 | -34.5176 | -20.9425
12.4344 | -34.518 | -20.9515
12.4405 | -34.5184 | -20.9605
12.4466 | -34.5189 | -20.9692
12.4527 | -34.5193 | -20.9783
12.4588 | -34.5198 | -20.9872
12.4649 | -34.5202 | -20.9963
12.471 | -34.5207 | -21.0058
12.4771 | -34.5211 | -21.0141
12.4832 | -34.5215 | -21.023
12.4893 | -34.522 | -21.0316
12.4954 | -34.5225 | -21.0404
12.5015 | -34.523 | -21.049
12.5076 | -34.5234 | -21.058
12.5137 | -34.5238 | -21.0673
12.5198 | -34.5242 | -21.076
12.5259 | -34.5247 | -21.0847
12.532 | -34.5252 | -21.0937
12.5382 | -34.5256 | -21.1025
12.5443 | -34.526 | -21.1115
12.5504 | -34.5265 | -21.1204
12.5565 | -34.527 | -21.1291
12.5626 | -34.5274 | -21.1382
12.5687 | -34.5279 | -21.1472
12.5748 | -34.5283 | -21.1561
12.5809 | -34.5287 | -21.1646
12.587 | -34.5291 | -21.174
12.5931 | -34.5296 | -21.1829
12.5992 | -34.53 | -21.1919
12.6053 | -34.5305 | -21.2007
12.6114 | -34.531 | -21.2095
12.6175 | -34.5314 | -21.2186
12.6236 | -34.5318 | -21.2272
12.6297 | -34.5323 | -21.2358
12.6358 | -34.5328 | -21.2445
12.6419 | -34.5333 | -21.2536
12.648 | -34.5337 | -21.2628
12.6541 | -34.5341 | -21.2715
12.6602 | -34.5346 | -21.2799
12.6663 | -34.535 | -21.2887
12.6724 | -34.5355 | -21.2976
12.6785 | -34.5359 | -21.3068
12.6847 | -34.5363 | -21.3157
12.6908 | -34.5368 | -21.3245
12.6969 | -34.5373 | -21.3333
12.703 | -34.5377 | -21.3423
12.7091 | -34.5382 | -21.3511
12.7152 | -34.5386 | -21.36
12.7213 | -34.539 | -21.369
12.7274 | -34.5395 | -21.3779
12.7335 | -34.54 | -21.3862
12.7396 | -34.5404 | -21.3946
12.7457 | -34.5408 | -21.4035
12.7518 | -34.5413 | -21.4123
12.7579 | -34.5417 | -21.4215
12.764 | -34.5421 | -21.4304
12.7701 | -34.5425 | -21.4389
12.7762 | -34.5429 | -21.4479
12.7823 | -34.5434 | -21.4567
12.7884 | -34.5438 | -21.4656
12.7945 | -34.5442 | -21.4741
12.8006 | -34.5447 | -21.4822
12.8067 | -34.5451 | -21.4913
12.8128 | -34.5456 | -21.4998
12.8189 | -34.5459 | -21.5082
12.8251 | -34.5463 | -21.5169
12.8312 | -34.5467 | -21.5257
12.8373 | -34.5472 | -21.5344
12.8434 | -34.5477 | -21.5432
12.8495 | -34.5481 | -21.5518
12.8556 | -34.5485 | -21.5605
12.8617 | -34.5489 | -21.5691
12.8678 | -34.5494 | -21.5779
12.8739 | -34.5498 | -21.5862
12.88 | -34.5502 | -21.595
12.8861 | -34.5506 | -21.6038
12.8922 | -34.5511 | -21.6125
12.8983 | -34.5516 | -21.6209
12.9044 | -34.552 | -21.6295
12.9105 | -34.5525 | -21.6382
12.9166 | -34.553 | -21.6473
12.9227 | -34.5534 | -21.6564
12.9288 | -34.5539 | -21.6653
12.9349 | -34.5544 | -21.6743
12.941 | -34.5549 | -21.6831
12.9471 | -34.5553 | -21.6918
12.9532 | -34.5558 | -21.7005
12.9593 | -34.5563 | -21.7095
12.9654 | -34.5567 | -21.7184
12.9716 | -34.5571 | -21.727
12.9777 | -34.5576 | -21.7354
12.9838 | -34.558 | -21.7441
12.9899 | -34.5584 | -21.7529
12.996 | -34.5588 | -21.7617
13.0021 | -34.5592 | -21.7704
13.0082 | -34.5597 | -21.7793
13.0143 | -34.5601 | -21.7875
13.0204 | -34.5605 | -21.7958
13.0265 | -34.561 | -21.8042
13.0326 | -34.5614 | -21.8131
13.0387 | -34.5619 | -21.822
13.0448 | -34.5623 | -21.8306
13.0509 | -34.5628 | -21.8391
13.057 | -34.5632 | -21.8473
13.0631 | -34.5637 | -21.8559
13.0692 | -34.5642 | -21.8645
13.0753 | -34.5647 | -21.8732
13.0814 | -34.5651 | -21.8816
13.0875 | -34.5655 | -21.89
13.0936 | -34.566 | -21.8986
13.0997 | -34.5664 | -21.9068
13.1058 | -34.5669 | -21.9149
13.112 | -34.5672 | -21.9236
13.1181 | -34.5677 | -21.9322
13.1242 | -34.5681 | -21.9408
13.1303 | -34.5686 | -21.9496
13.1364 | -34.569 | -21.9582
13.1425 | -34.5694 | -21.9668
13.1486 | -34.5699 | -21.9752
13.1547 | -34.5703 | -21.9835
13.1608 | -34.5708 | -21.992
13.1669 | -34.5712 | -22.0004
13.173 | -34.5717 | -22.009
13.1791 | -34.5721 | -22.0177
13.1852 | -34.5725 | -22.0265
13.1913 | -34.5729 | -22.0352
13.1974 | -34.5733 | -22.0437
13.2035 | -34.5738 | -22.0523
13.2096 | -34.5743 | -22.0609
13.2157 | -34.5747 | -22.0695
13.2218 | -34.5752 | -22.0782
13.2279 | -34.5756 | -22.0866
13.234 | -34.5761 | -22.0951
13.2401 | -34.5765 | -22.1036
13.2462 | -34.5769 | -22.1121
13.2524 | -34.5773 | -22.1205
13.2585 | -34.5778 | -22.129
13.2646 | -34.5782 | -22.1377
13.2707 | -34.5787 | -22.1462
13.2768 | -34.5791 | -22.1551
13.2829 | -34.5795 | -22.1639
13.289 | -34.58 | -22.1726
13.2951 | -34.5804 | -22.1808
13.3012 | -34.5808 | -22.1895
13.3073 | -34.5813 | -22.1983
13.3134 | -34.5817 | -22.2068
13.3195 | -34.5821 | -22.2153
13.3256 | -34.5826 | -22.224
13.3317 | -34.583 | -22.2325
13.3378 | -34.5835 | -22.2411
13.3439 | -34.5839 | -22.2496
13.35 | -34.5843 | -22.2581
13.3561 | -34.5847 | -22.2668
13.3622 | -34.5852 | -22.2753
13.3683 | -34.5857 | -22.2836
13.3744 | -34.5861 | -22.2921
13.3805 | -34.5865 | -22.3006
13.3866 | -34.5869 | -22.3089
13.3927 | -34.5873 | -22.3172
13.3989 | -34.5878 | -22.3254
13.405 | -34.5882 | -22.3338
13.4111 | -34.5887 | -22.3423
13.4172 | -34.5891 | -22.3505
13.4233 | -34.5895 | -22.3593
13.4294 | -34.59 | -22.3675
13.4355 | -34.5904 | -22.3766
13.4416 | -34.5908 | -22.3851
13.4477 | -34.5912 | -22.3937
13.4538 | -34.5916 | -22.4021
13.4599 | -34.5921 | -22.4109
13.466 | -34.5926 | -22.4197
13.4721 | -34.593 | -22.4282
13.4782 | -34.5935 | -22.4366
13.4843 | -34.594 | -22.4452
13.4904 | -34.5944 | -22.4532
13.4965 | -34.5948 | -22.4621
13.5026 | -34.5952 | -22.4706
13.5087 | -34.5957 | -22.4787
13.5148 | -34.5961 | -22.4873
13.5209 | -34.5966 | -22.4955
13.527 | -34.597 | -22.5043
13.5331 | -34.5974 | -22.5126
13.5393 | -34.5979 | -22.521
13.5454 | -34.5983 | -22.5295
13.5515 | -34.5988 | -22.5379
13.5576 | -34.5992 | -22.5463
13.5637 | -34.5997 | -22.5547
13.5698 | -34.6001 | -22.5634
13.5759 | -34.6005 | -22.5715
13.582 | -34.6009 | -22.5801
13.5881 | -34.6013 | -22.5886
13.5942 | -34.6018 | -22.5972
13.6003 | -34.6022 | -22.6057
13.6064 | -34.6026 | -22.6143
13.6125 | -34.6031 | -22.623
13.6186 | -34.6036 | -22.6315
13.6247 | -34.604 | -22.6399
13.6308 | -34.6044 | -22.6481
13.6369 | -34.6049 | -22.6566
13.643 | -34.6053 | -22.665
13.6491 | -34.6058 | -22.6734
13.6552 | -34.6062 | -22.6816
13.6613 | -34.6066 | -22.6902
13.6674 | -34.6071 | -22.6988
13.6735 | -34.6075 | -22.7074
13.6796 | -34.6079 | -22.7156
13.6858 | -34.6083 | -22.724
13.6919 | -34.6088 | -22.7324
13.698 | -34.6092 | -22.741
13.7041 | -34.6096 | -22.7497
13.7102 | -34.6101 | -22.7583
13.7163 | -34.6106 | -22.7667
13.7224 | -34.6111 | -22.7749
13.7285 | -34.6115 | -22.7836
13.7346 | -34.6119 | -22.7921
13.7407 | -34.6124 | -22.8006
13.7468 | -34.6128 | -22.8091
13.7529 | -34.6132 | -22.8173
13.759 | -34.6137 | -22.8259
13.7651 | -34.6141 | -22.8345
13.7712 | -34.6145 | -22.8431
13.7773 | -34.615 | -22.8517
13.7834 | -34.6154 | -22.86
13.7895 | -34.616 | -22.8683
13.7956 | -34.6164 | -22.8769
13.8017 | -34.6168 | -22.8853
13.8078 | -34.6173 | -22.8937
13.8139 | -34.6177 | -22.9019
13.82 | -34.6182 | -22.9102
13.8262 | -34.6186 | -22.9189
13.8323 | -34.6191 | -22.9276
13.8384 | -34.6195 | -22.9362
13.8445 | -34.62 | -22.9448
13.8506 | -34.6204 | -22.9534
13.8567 | -34.6209 | -22.962
13.8628 | -34.6213 | -22.9705
13.8689 | -34.6217 | -22.9791
13.875 | -34.6222 | -22.9874
13.8811 | -34.6226 | -22.9958
13.8872 | -34.6231 | -23.0042
13.8933 | -34.6236 | -23.0126
13.8994 | -34.624 | -23.0207
13.9055 | -34.6245 | -23.0292
13.9116 | -34.6249 | -23.0373
13.9177 | -34.6254 | -23.0454
13.9238 | -34.6259 | -23.0541
13.9299 | -34.6264 | -23.0623
13.936 | -34.6268 | -23.0706
13.9421 | -34.6272 | -23.0791
13.9482 | -34.6277 | -23.0876
13.9543 | -34.6281 | -23.0958
13.9604 | -34.6286 | -23.1044
13.9665 | -34.629 | -23.1128
13.9727 | -34.6294 | -23.121
13.9788 | -34.6299 | -23.1295
13.9849 | -34.6303 | -23.1381
13.991 | -34.6307 | -23.1467
13.9971 | -34.6311 | -23.1554
14.0032 | -34.6315 | -23.164
14.0093 | -34.6319 | -23.1723
14.0154 | -34.6324 | -23.181
14.0215 | -34.6328 | -23.1892
14.0276 | -34.6333 | -23.1974
14.0337 | -34.6337 | -23.2061
14.0398 | -34.6342 | -23.2143
14.0459 | -34.6346 | -23.2228
14.052 | -34.635 | -23.2313
14.0581 | -34.6355 | -23.2393
14.0642 | -34.6359 | -23.2479
14.0703 | -34.6364 | -23.2561
14.0764 | -34.6368 | -23.2649
14.0825 | -34.6373 | -23.2735
14.0886 | -34.6377 | -23.2814
14.0947 | -34.6381 | -23.2895
14.1008 | -34.6386 | -23.2983
14.1069 | -34.639 | -23.3065
14.1131 | -34.6394 | -23.3152
14.1192 | -34.6398 | -23.3233
14.1253 | -34.6402 | -23.3314
14.1314 | -34.6407 | -23.3398
14.1375 | -34.6412 | -23.3484
14.1436 | -34.6417 | -23.3567
14.1497 | -34.6421 | -23.3648
14.1558 | -34.6426 | -23.3734
14.1619 | -34.643 | -23.3817
14.168 | -34.6435 | -23.3899
14.1741 | -34.644 | -23.3982
14.1802 | -34.6444 | -23.4064
14.1863 | -34.6449 | -23.4148
14.1924 | -34.6453 | -23.4233
14.1985 | -34.6458 | -23.4314
14.2046 | -34.6462 | -23.4396
14.2107 | -34.6467 | -23.4475
14.2168 | -34.6472 | -23.4563
14.2229 | -34.6476 | -23.4647
14.229 | -34.6481 | -23.4731
14.2351 | -34.6485 | -23.4812
14.2412 | -34.649 | -23.4894
14.2473 | -34.6495 | -23.4978
14.2534 | -34.6499 | -23.506
14.2596 | -34.6504 | -23.5144
14.2657 | -34.6508 | -23.5228
14.2718 | -34.6513 | -23.5313
14.2779 | -34.6517 | -23.5394
14.284 | -34.6522 | -23.5477
14.2901 | -34.6527 | -23.5558
14.2962 | -34.6532 | -23.5639
14.3023 | -34.6536 | -23.5723
14.3084 | -34.6541 | -23.5809
14.3145 | -34.6545 | -23.5891
14.3206 | -34.6549 | -23.5975
14.3267 | -34.6554 | -23.6057
14.3328 | -34.6559 | -23.614
14.3389 | -34.6563 | -23.622
14.345 | -34.6567 | -23.6304
14.3511 | -34.6572 | -23.6384
14.3572 | -34.6576 | -23.647
14.3633 | -34.6581 | -23.6552
14.3694 | -34.6585 | -23.6634
14.3755 | -34.6589 | -23.6714
14.3816 | -34.6594 | -23.6796
14.3877 | -34.6599 | -23.6877
14.3938 | -34.6603 | -23.6962
14.4 | -34.6608 | -23.7043
14.4061 | -34.6612 | -23.7126
14.4122 | -34.6616 | -23.721
14.4183 | -34.662 | -23.7293
14.4244 | -34.6624 | -23.7374
14.4305 | -34.6629 | -23.7459
14.4366 | -34.6634 | -23.7543
14.4427 | -34.6638 | -23.7629
14.4488 | -34.6642 | -23.7707
14.4549 | -34.6647 | -23.7789
14.461 | -34.6652 | -23.7873
14.4671 | -34.6656 | -23.7956
14.4732 | -34.6661 | -23.8036
14.4793 | -34.6665 | -23.8125
14.4854 | -34.667 | -23.8204
14.4915 | -34.6675 | -23.8283
14.4976 | -34.668 | -23.8368
14.5037 | -34.6685 | -23.8451
14.5098 | -34.6689 | -23.8532
14.5159 | -34.6694 | -23.8617
14.522 | -34.6698 | -23.8701
14.5281 | -34.6703 | -23.878
14.5342 | -34.6707 | -23.8865
14.5403 | -34.6712 | -23.8948
14.5465 | -34.6717 | -23.9034
14.5526 | -34.6721 | -23.912
14.5587 | -34.6726 | -23.9202
14.5648 | -34.673 | -23.9282
14.5709 | -34.6735 | -23.9367
14.577 | -34.674 | -23.9449
14.5831 | -34.6744 | -23.9533
14.5892 | -34.675 | -23.9611
14.5953 | -34.6754 | -23.9691
14.6014 | -34.6759 | -23.9773
14.6075 | -34.6763 | -23.9852
14.6136 | -34.6768 | -23.9932
14.6197 | -34.6772 | -24.0016
14.6258 | -34.6777 | -24.01
14.6319 | -34.6782 | -24.018
14.638 | -34.6787 | -24.0262
14.6441 | -34.6791 | -24.0339
14.6502 | -34.6795 | -24.0419
14.6563 | -34.68 | -24.05
14.6624 | -34.6805 | -24.0581
14.6685 | -34.681 | -24.066
14.6746 | -34.6814 | -24.0742
14.6807 | -34.6819 | -24.0824
14.6869 | -34.6823 | -24.0905
14.693 | -34.6828 | -24.0986
14.6991 | -34.6832 | -24.1066
14.7052 | -34.6837 | -24.1147
14.7113 | -34.6842 | -24.1231
14.7174 | -34.6846 | -24.1312
14.7235 | -34.6851 | -24.1392
14.7296 | -34.6855 | -24.1472
14.7357 | -34.6861 | -24.1552
14.7418 | -34.6865 | -24.1631
14.7479 | -34.687 | -24.1713
14.754 | -34.6875 | -24.1794
14.7601 | -34.6879 | -24.1877
14.7662 | -34.6884 | -24.1959
14.7723 | -34.6889 | -24.2039
14.7784 | -34.6893 | -24.2121
14.7845 | -34.6898 | -24.2203
14.7906 | -34.6903 | -24.2285
14.7967 | -34.6907 | -24.2366
14.8028 | -34.6911 | -24.2445
14.8089 | -34.6916 | -24.2525
14.815 | -34.6921 | -24.2604
14.8211 | -34.6925 | -24.2685
14.8272 | -34.6929 | -24.2767
14.8334 | -34.6934 | -24.2847
14.8395 | -34.6939 | -24.2929
14.8456 | -34.6944 | -24.3009
14.8517 | -34.6949 | -24.3091
14.8578 | -34.6953 | -24.3171
14.8639 | -34.6958 | -24.3255
14.87 | -34.6963 | -24.3337
14.8761 | -34.6968 | -24.3415
14.8822 | -34.6972 | -24.3497
14.8883 | -34.6977 | -24.3576
14.8944 | -34.6982 | -24.3657
14.9005 | -34.6987 | -24.3737
14.9066 | -34.6991 | -24.3818
14.9127 | -34.6996 | -24.3897
14.9188 | -34.7001 | -24.3981
14.9249 | -34.7005 | -24.4062
14.931 | -34.701 | -24.4141
14.9371 | -34.7015 | -24.4221
14.9432 | -34.702 | -24.4303
14.9493 | -34.7024 | -24.4383
14.9554 | -34.7029 | -24.4463
14.9615 | -34.7033 | -24.4545
14.9676 | -34.7038 | -24.4627
14.9738 | -34.7043 | -24.4705
14.9799 | -34.7048 | -24.479
14.986 | -34.7053 | -24.4867
14.9921 | -34.7058 | -24.4947
14.9982 | -34.7062 | -24.503
15.0043 | -34.7066 | -24.511
15.0104 | -34.7071 | -24.5191
15.0165 | -34.7076 | -24.5272
15.0226 | -34.7081 | -24.5349
15.0287 | -34.7085 | -24.5434
15.0348 | -34.7091 | -24.5516
15.0409 | -34.7095 | -24.56
15.047 | -34.71 | -24.5684
15.0531 | -34.7106 | -24.5766
15.0592 | -34.7111 | -24.5849
15.0653 | -34.7115 | -24.5929
15.0714 | -34.712 | -24.6006
15.0775 | -34.7125 | -24.6084
15.0836 | -34.713 | -24.6166
15.0897 | -34.7135 | -24.6245
15.0958 | -34.714 | -24.6323
15.1019 | -34.7144 | -24.6404
15.108 | -34.7149 | -24.6485
15.1141 | -34.7153 | -24.6566
15.1203 | -34.7158 | -24.6648
15.1264 | -34.7164 | -24.6729
15.1325 | -34.7169 | -24.6809
15.1386 | -34.7173 | -24.689
15.1447 | -34.7179 | -24.697
15.1508 | -34.7184 | -24.7054
15.1569 | -34.7189 | -24.7137
15.163 | -34.7194 | -24.7212
15.1691 | -34.7199 | -24.7292
15.1752 | -34.7204 | -24.7372
15.1813 | -34.7209 | -24.7452
15.1874 | -34.7214 | -24.7534
15.1935 | -34.7219 | -24.7612
15.1996 | -34.7225 | -24.7693
15.2057 | -34.7229 | -24.7772
15.2118 | -34.7234 | -24.7853
15.2179 | -34.7239 | -24.7931
15.224 | -34.7243 | -24.8012
15.2301 | -34.7248 | -24.8089
15.2362 | -34.7253 | -24.8168
15.2423 | -34.7258 | -24.8251
15.2484 | -34.7263 | -24.8331
15.2545 | -34.7268 | -24.8411
15.2607 | -34.7273 | -24.8492
15.2668 | -34.7279 | -24.857
15.2729 | -34.7284 | -24.8648
15.279 | -34.7289 | -24.8729
15.2851 | -34.7294 | -24.8809
15.2912 | -34.7298 | -24.8888
15.2973 | -34.7303 | -24.8968
15.3034 | -34.7308 | -24.9046
15.3095 | -34.7313 | -24.9123
15.3156 | -34.7318 | -24.9204
15.3217 | -34.7323 | -24.9281
15.3278 | -34.7328 | -24.9361
15.3339 | -34.7333 | -24.9438
15.34 | -34.7338 | -24.9515
15.3461 | -34.7343 | -24.9595
15.3522 | -34.7348 | -24.9674
15.3583 | -34.7353 | -24.9753
15.3644 | -34.7358 | -24.9834
15.3705 | -34.7363 | -24.9911
15.3766 | -34.7368 | -24.9993
15.3827 | -34.7373 | -25.007
15.3888 | -34.7378 | -25.0147
15.3949 | -34.7383 | -25.0224
15.401 | -34.7388 | -25.0305
15.4072 | -34.7393 | -25.0383
15.4133 | -34.7398 | -25.0462
15.4194 | -34.7403 | -25.0542
15.4255 | -34.7408 | -25.0622
15.4316 | -34.7413 | -25.0701
15.4377 | -34.7418 | -25.078
15.4438 | -34.7423 | -25.0856
15.4499 | -34.7428 | -25.0935
15.456 | -34.7433 | -25.1013
15.4621 | -34.7438 | -25.1093
15.4682 | -34.7443 | -25.1173
15.4743 | -34.7449 | -25.1252
15.4804 | -34.7453 | -25.1331
15.4865 | -34.7458 | -25.141
15.4926 | -34.7463 | -25.1491
15.4987 | -34.7468 | -25.1571
15.5048 | -34.7473 | -25.1649
15.5109 | -34.7478 | -25.1725
15.517 | -34.7484 | -25.1805
15.5231 | -34.7489 | -25.1884
15.5292 | -34.7494 | -25.1961
15.5353 | -34.7499 | -25.2039
15.5414 | -34.7504 | -25.212
15.5476 | -34.7509 | -25.22
15.5537 | -34.7514 | -25.2279
15.5598 | -34.7519 | -25.2355
15.5659 | -34.7524 | -25.2436
15.572 | -34.7529 | -25.2512
15.5781 | -34.7534 | -25.2591
15.5842 | -34.754 | -25.2667
15.5903 | -34.7545 | -25.2743
15.5964 | -34.755 | -25.282
15.6025 | -34.7555 | -25.2897
15.6086 | -34.756 | -25.2977
15.6147 | -34.7565 | -25.3056
15.6208 | -34.757 | -25.3133
15.6269 | -34.7576 | -25.3211
15.633 | -34.758 | -25.329
15.6391 | -34.7585 | -25.337
15.6452 | -34.759 | -25.3447
15.6513 | -34.7595 | -25.3527
15.6574 | -34.7601 | -25.3603
15.6635 | -34.7605 | -25.3682
15.6696 | -34.7611 | -25.376
15.6757 | -34.7616 | -25.3842
15.6818 | -34.7621 | -25.3923
15.688 | -34.7627 | -25.4002
15.6941 | -34.7631 | -25.408
15.7002 | -34.7636 | -25.4158
15.7063 | -34.7641 | -25.4238
15.7124 | -34.7646 | -25.4318
15.7185 | -34.7651 | -25.4396
15.7246 | -34.7656 | -25.4472
15.7307 | -34.7662 | -25.4552
15.7368 | -34.7667 | -25.4633
15.7429 | -34.7672 | -25.4711
15.749 | -34.7677 | -25.4789
15.7551 | -34.7682 | -25.4868
15.7612 | -34.7687 | -25.494
15.7673 | -34.7692 | -25.5017
15.7734 | -34.7696 | -25.5093
15.7795 | -34.7702 | -25.5176
15.7856 | -34.7707 | -25.5252
15.7917 | -34.7712 | -25.5329
15.7978 | -34.7717 | -25.5406
15.8039 | -34.7722 | -25.5483
15.81 | -34.7727 | -25.5561
15.8161 | -34.7732 | -25.564
15.8222 | -34.7737 | -25.5718
15.8283 | -34.7742 | -25.5795
15.8345 | -34.7747 | -25.5874
15.8406 | -34.7753 | -25.5952
15.8467 | -34.7758 | -25.603
15.8528 | -34.7763 | -25.6108
15.8589 | -34.7768 | -25.6185
15.865 | -34.7773 | -25.6262
15.8711 | -34.7778 | -25.6343
15.8772 | -34.7784 | -25.642
15.8833 | -34.7789 | -25.6501
15.8894 | -34.7794 | -25.6578
15.8955 | -34.7799 | -25.6658
15.9016 | -34.7804 | -25.6735
15.9077 | -34.7809 | -25.6813
15.9138 | -34.7814 | -25.6889
15.9199 | -34.7819 | -25.6969
15.926 | -34.7824 | -25.7049
15.9321 | -34.783 | -25.7126
15.9382 | -34.7836 | -25.7202
15.9443 | -34.7841 | -25.7283
15.9504 | -34.7846 | -25.7362
15.9565 | -34.7852 | -25.744
15.9626 | -34.7857 | -25.7518
15.9687 | -34.7862 | -25.7594
15.9749 | -34.7867 | -25.7674
15.981 | -34.7873 | -25.7753
15.9871 | -34.7878 | -25.783
15.9932 | -34.7883 | -25.791
15.9993 | -34.7888 | -25.7985
16.0054 | -34.7893 | -25.806
16.0115 | -34.7899 | -25.8137
16.0176 | -34.7904 | -25.8218
16.0237 | -34.7909 | -25.8298
16.0298 | -34.7914 | -25.8371
16.0359 | -34.7919 | -25.8447
16.042 | -34.7925 | -25.8524
16.0481 | -34.793 | -25.8602
16.0542 | -34.7935 | -25.8683
16.0603 | -34.7941 | -25.8759
16.0664 | -34.7946 | -25.8837
16.0725 | -34.7951 | -25.8913
16.0786 | -34.7957 | -25.8994
16.0847 | -34.7962 | -25.9069
16.0908 | -34.7967 | -25.9145
16.0969 | -34.7972 | -25.922
16.103 | -34.7977 | -25.93
16.1091 | -34.7982 | -25.9377
16.1152 | -34.7987 | -25.9456
16.1214 | -34.7992 | -25.9534
16.1275 | -34.7997 | -25.961
16.1336 | -34.8003 | -25.9687
16.1397 | -34.8008 | -25.9765
16.1458 | -34.8013 | -25.9841
16.1519 | -34.8019 | -25.9921
16.158 | -34.8024 | -25.9999
16.1641 | -34.8029 | -26.0078
16.1702 | -34.8034 | -26.0156
16.1763 | -34.8039 | -26.0235
16.1824 | -34.8045 | -26.031
16.1885 | -34.805 | -26.0387
16.1946 | -34.8055 | -26.0463
16.2007 | -34.8061 | -26.0544
16.2068 | -34.8066 | -26.0619
16.2129 | -34.8071 | -26.0698
16.219 | -34.8076 | -26.0774
16.2251 | -34.8081 | -26.0853
16.2312 | -34.8086 | -26.0932
16.2373 | -34.8091 | -26.1007
16.2434 | -34.8096 | -26.1084
16.2495 | -34.8101 | -26.116
16.2556 | -34.8107 | -26.1241
16.2618 | -34.8112 | -26.1315
16.2679 | -34.8117 | -26.1393
16.274 | -34.8123 | -26.1471
16.2801 | -34.8128 | -26.1547
16.2862 | -34.8133 | -26.1623
16.2923 | -34.8138 | -26.1703
16.2984 | -34.8143 | -26.1782
16.3045 | -34.8149 | -26.1859
16.3106 | -34.8155 | -26.1934
16.3167 | -34.8159 | -26.2015
16.3228 | -34.8165 | -26.2092
16.3289 | -34.817 | -26.2169
16.335 | -34.8175 | -26.2248
16.3411 | -34.818 | -26.2326
16.3472 | -34.8185 | -26.2407
16.3533 | -34.819 | -26.2484
16.3594 | -34.8195 | -26.2559
16.3655 | -34.82 | -26.2635
16.3716 | -34.8205 | -26.2711
16.3777 | -34.821 | -26.2788
16.3838 | -34.8215 | -26.2863
16.3899 | -34.822 | -26.2942
16.396 | -34.8225 | -26.3021
16.4021 | -34.823 | -26.3096
16.4083 | -34.8236 | -26.3173
16.4144 | -34.8242 | -26.325
16.4205 | -34.8247 | -26.3328
16.4266 | -34.8252 | -26.3404
16.4327 | -34.8257 | -26.3479
16.4388 | -34.8262 | -26.3557
16.4449 | -34.8267 | -26.3634
16.451 | -34.8272 | -26.3714
16.4571 | -34.8277 | -26.3793
16.4632 | -34.8282 | -26.3868
16.4693 | -34.8287 | -26.3944
16.4754 | -34.8292 | -26.4023
16.4815 | -34.8298 | -26.4102
16.4876 | -34.8303 | -26.4177
16.4937 | -34.8308 | -26.4254
16.4998 | -34.8314 | -26.4332
16.5059 | -34.8319 | -26.4408
16.512 | -34.8324 | -26.4487
16.5181 | -34.833 | -26.4565
16.5242 | -34.8335 | -26.4643
16.5303 | -34.834 | -26.4717
16.5364 | -34.8345 | -26.4793
16.5425 | -34.835 | -26.4871
16.5487 | -34.8355 | -26.4946
16.5548 | -34.836 | -26.5022
16.5609 | -34.8365 | -26.5101
16.567 | -34.837 | -26.5179
16.5731 | -34.8376 | -26.5255
16.5792 | -34.8381 | -26.533
16.5853 | -34.8386 | -26.5405
16.5914 | -34.8391 | -26.5482
16.5975 | -34.8396 | -26.5561
16.6036 | -34.8401 | -26.5637
16.6097 | -34.8406 | -26.5713
16.6158 | -34.8411 | -26.579
16.6219 | -34.8417 | -26.5865
16.628 | -34.8421 | -26.5944
16.6341 | -34.8427 | -26.602
16.6402 | -34.8432 | -26.6093
16.6463 | -34.8437 | -26.6169
16.6524 | -34.8442 | -26.6243
16.6585 | -34.8448 | -26.6321
16.6646 | -34.8453 | -26.6398
16.6707 | -34.8458 | -26.6471
16.6768 | -34.8464 | -26.6547
16.6829 | -34.8469 | -26.6624
16.689 | -34.8474 | -26.6701
16.6952 | -34.8479 | -26.678
16.7013 | -34.8484 | -26.6856
16.7074 | -34.8489 | -26.6934
16.7135 | -34.8494 | -26.7011
16.7196 | -34.85 | -26.7088
16.7257 | -34.8505 | -26.7164
16.7318 | -34.851 | -26.724
16.7379 | -34.8515 | -26.7318
16.744 | -34.852 | -26.7393
16.7501 | -34.8525 | -26.7469
16.7562 | -34.8531 | -26.7546
16.7623 | -34.8536 | -26.7623
16.7684 | -34.8541 | -26.7701
16.7745 | -34.8546 | -26.7774
16.7806 | -34.8551 | -26.7852
16.7867 | -34.8556 | -26.7929
16.7928 | -34.8561 | -26.8008
16.7989 | -34.8566 | -26.8082
16.805 | -34.8571 | -26.816
16.8111 | -34.8576 | -26.8236
16.8172 | -34.8581 | -26.8318
16.8233 | -34.8586 | -26.8395
16.8294 | -34.8591 | -26.8473
16.8356 | -34.8596 | -26.8551
16.8417 | -34.8602 | -26.8627
16.8478 | -34.8606 | -26.8702
16.8539 | -34.8612 | -26.8778
16.86 | -34.8616 | -26.8854
16.8661 | -34.8621 | -26.8929
16.8722 | -34.8627 | -26.9006
16.8783 | -34.8632 | -26.908
16.8844 | -34.8637 | -26.9157
16.8905 | -34.8643 | -26.9235
16.8966 | -34.8647 | -26.9311
16.9027 | -34.8653 | -26.9386
16.9088 | -34.8658 | -26.9467
16.9149 | -34.8663 | -26.9548
16.921 | -34.8668 | -26.9623
16.9271 | -34.8673 | -26.9703
16.9332 | -34.8678 | -26.978
16.9393 | -34.8683 | -26.9859
16.9454 | -34.8688 | -26.9933
16.9515 | -34.8693 | -27.0013
16.9576 | -34.8698 | -27.009
16.9637 | -34.8704 | -27.0166
16.9698 | -34.8709 | -27.024
16.9759 | -34.8714 | -27.0315
16.9821 | -34.8719 | -27.0393
16.9882 | -34.8724 | -27.047
16.9943 | -34.8729 | -27.0549
17.0004 | -34.8734 | -27.0626
17.0065 | -34.8739 | -27.0703
17.0126 | -34.8744 | -27.0779
17.0187 | -34.8749 | -27.0857
17.0248 | -34.8755 | -27.093
17.0309 | -34.876 | -27.1006
17.037 | -34.8765 | -27.1083
17.0431 | -34.877 | -27.1156
17.0492 | -34.8775 | -27.1233
17.0553 | -34.878 | -27.131
17.0614 | -34.8784 | -27.1385
17.0675 | -34.879 | -27.1464
17.0736 | -34.8795 | -27.1541
17.0797 | -34.8799 | -27.1613
17.0858 | -34.8804 | -27.1693
17.0919 | -34.8809 | -27.1771
17.098 | -34.8814 | -27.1845
17.1041 | -34.882 | -27.1921
17.1102 | -34.8825 | -27.1999
17.1163 | -34.883 | -27.2077
17.1225 | -34.8835 | -27.2151
17.1286 | -34.884 | -27.2225
17.1347 | -34.8845 | -27.23
17.1408 | -34.885 | -27.2376
17.1469 | -34.8856 | -27.2451
17.153 | -34.8861 | -27.2527
17.1591 | -34.8866 | -27.26
17.1652 | -34.8871 | -27.2674
17.1713 | -34.8876 | -27.2749
17.1774 | -34.8881 | -27.2826
17.1835 | -34.8886 | -27.2904
17.1896 | -34.8891 | -27.2979
17.1957 | -34.8896 | -27.3055
17.2018 | -34.8901 | -27.3132
17.2079 | -34.8906 | -27.3207
17.214 | -34.8911 | -27.3282
17.2201 | -34.8916 | -27.3358
17.2262 | -34.8921 | -27.3435
17.2323 | -34.8926 | -27.3511
17.2384 | -34.8931 | -27.3587
17.2445 | -34.8936 | -27.3662
17.2506 | -34.8941 | -27.3739
17.2567 | -34.8946 | -27.3815
17.2628 | -34.8951 | -27.3893
17.269 | -34.8955 | -27.397
17.2751 | -34.8961 | -27.4047
17.2812 | -34.8966 | -27.4123
17.2873 | -34.8971 | -27.42
17.2934 | -34.8975 | -27.4277
17.2995 | -34.898 | -27.4352
17.3056 | -34.8985 | -27.4427
17.3117 | -34.899 | -27.4502
17.3178 | -34.8995 | -27.4578
17.3239 | -34.9 | -27.4654
17.33 | -34.9005 | -27.4728
17.3361 | -34.9009 | -27.4804
17.3422 | -34.9014 | -27.4879
17.3483 | -34.9019 | -27.4955
17.3544 | -34.9024 | -27.5031
17.3605 | -34.9029 | -27.5107
17.3666 | -34.9034 | -27.5184
17.3727 | -34.9039 | -27.5259
17.3788 | -34.9044 | -27.5333
17.3849 | -34.9049 | -27.5408
17.391 | -34.9054 | -27.5482
17.3971 | -34.906 | -27.5558
17.4032 | -34.9064 | -27.5637
17.4094 | -34.9069 | -27.5713
17.4155 | -34.9074 | -27.5787
17.4216 | -34.9079 | -27.5863
17.4277 | -34.9084 | -27.5941
17.4338 | -34.9089 | -27.6018
17.4399 | -34.9094 | -27.6095
17.446 | -34.9099 | -27.6169
17.4521 | -34.9104 | -27.6243
17.4582 | -34.9109 | -27.6316
17.4643 | -34.9113 | -27.6393
17.4704 | -34.9118 | -27.6468
17.4765 | -34.9123 | -27.6544
17.4826 | -34.9128 | -27.6619
17.4887 | -34.9133 | -27.6696
17.4948 | -34.9138 | -27.677
17.5009 | -34.9143 | -27.6844
17.507 | -34.9148 | -27.6921
17.5131 | -34.9153 | -27.6997
17.5192 | -34.9157 | -27.707
17.5253 | -34.9162 | -27.7147
17.5314 | -34.9166 | -27.7222
17.5375 | -34.9171 | -27.7299
17.5436 | -34.9176 | -27.7376
17.5497 | -34.9181 | -27.7454
17.5559 | -34.9186 | -27.7527
17.562 | -34.9191 | -27.7602
17.5681 | -34.9196 | -27.7676
17.5742 | -34.92 | -27.7752
17.5803 | -34.9205 | -27.7826
17.5864 | -34.9211 | -27.7899
17.5925 | -34.9216 | -27.7978
17.5986 | -34.9221 | -27.8051
17.6047 | -34.9226 | -27.813
17.6108 | -34.9231 | -27.8201
17.6169 | -34.9236 | -27.8277
17.623 | -34.9241 | -27.8351
17.6291 | -34.9246 | -27.8426
17.6352 | -34.9251 | -27.85
17.6413 | -34.9256 | -27.8577
17.6474 | -34.9261 | -27.8654
17.6535 | -34.9266 | -27.873
17.6596 | -34.927 | -27.8805
17.6657 | -34.9275 | -27.8879
17.6718 | -34.928 | -27.8955
17.6779 | -34.9285 | -27.9029
17.684 | -34.929 | -27.9105
17.6901 | -34.9295 | -27.9178
17.6963 | -34.93 | -27.9251
17.7024 | -34.9305 | -27.9326
17.7085 | -34.9309 | -27.9398
17.7146 | -34.9314 | -27.9474
17.7207 | -34.9319 | -27.9547
17.7268 | -34.9324 | -27.9621
17.7329 | -34.9329 | -27.9701
17.739 | -34.9334 | -27.9777
17.7451 | -34.9339 | -27.9851
17.7512 | -34.9344 | -27.9928
17.7573 | -34.9348 | -28.0004
17.7634 | -34.9354 | -28.0076
17.7695 | -34.9358 | -28.0152
17.7756 | -34.9363 | -28.0226
17.7817 | -34.9368 | -28.0298
17.7878 | -34.9373 | -28.0373
17.7939 | -34.9378 | -28.0447
17.8 | -34.9383 | -28.052
17.8061 | -34.9388 | -28.0596
17.8122 | -34.9392 | -28.0671
17.8183 | -34.9397 | -28.0747
17.8244 | -34.9402 | -28.0822
17.8305 | -34.9407 | -28.0898
17.8366 | -34.9412 | -28.0969
17.8428 | -34.9417 | -28.1046
17.8489 | -34.9422 | -28.1119
17.855 | -34.9427 | -28.1192
17.8611 | -34.9432 | -28.1265
17.8672 | -34.9437 | -28.1338
17.8733 | -34.9442 | -28.1416
17.8794 | -34.9447 | -28.1492
17.8855 | -34.9452 | -28.1566
17.8916 | -34.9457 | -28.1639
17.8977 | -34.9462 | -28.1712
17.9038 | -34.9466 | -28.1785
17.9099 | -34.9471 | -28.1859
17.916 | -34.9476 | -28.1934
17.9221 | -34.9481 | -28.2009
17.9282 | -34.9485 | -28.2082
17.9343 | -34.949 | -28.2157
17.9404 | -34.9495 | -28.223
17.9465 | -34.95 | -28.2305
17.9526 | -34.9504 | -28.2382
17.9587 | -34.9509 | -28.2456
17.9648 | -34.9514 | -28.253
17.9709 | -34.9519 | -28.2604
17.977 | -34.9524 | -28.2677
17.9832 | -34.9529 | -28.2752
17.9893 | -34.9533 | -28.2829
17.9954 | -34.9538 | -28.2906
18.0015 | -34.9544 | -28.2981
18.0076 | -34.9548 | -28.3051
18.0137 | -34.9553 | -28.3128
18.0198 | -34.9558 | -28.3202
18.0259 | -34.9562 | -28.3277
18.032 | -34.9567 | -28.335
18.0381 | -34.9572 | -28.3423
18.0442 | -34.9577 | -28.3498
18.0503 | -34.9581 | -28.3573
18.0564 | -34.9586 | -28.3647
18.0625 | -34.9591 | -28.3723
18.0686 | -34.9596 | -28.3795
18.0747 | -34.9601 | -28.387
18.0808 | -34.9605 | -28.3941
18.0869 | -34.961 | -28.4016
18.093 | -34.9616 | -28.409
18.0991 | -34.962 | -28.4166
18.1052 | -34.9625 | -28.4243
18.1113 | -34.963 | -28.4314
18.1174 | -34.9635 | -28.4386
18.1236 | -34.964 | -28.4461
18.1297 | -34.9645 | -28.4534
18.1358 | -34.965 | -28.4608
18.1419 | -34.9655 | -28.4683
18.148 | -34.966 | -28.4757
18.1541 | -34.9665 | -28.4829
18.1602 | -34.967 | -28.4905
18.1663 | -34.9676 | -28.4978
18.1724 | -34.9681 | -28.5054
18.1785 | -34.9685 | -28.5128
18.1846 | -34.969 | -28.5201
18.1907 | -34.9695 | -28.5279
18.1968 | -34.9699 | -28.5354
18.2029 | -34.9704 | -28.5424
18.209 | -34.9709 | -28.5501
18.2151 | -34.9714 | -28.5578
18.2212 | -34.9719 | -28.5649
18.2273 | -34.9724 | -28.5723
18.2334 | -34.9729 | -28.5797
18.2395 | -34.9734 | -28.5869
18.2456 | -34.9739 | -28.5943
18.2517 | -34.9743 | -28.6017
18.2578 | -34.9748 | -28.6089
18.2639 | -34.9753 | -28.6166
18.2701 | -34.9758 | -28.6239
18.2762 | -34.9763 | -28.6313
18.2823 | -34.9768 | -28.6388
18.2884 | -34.9773 | -28.6463
18.2945 | -34.9778 | -28.6534
18.3006 | -34.9783 | -28.6608
18.3067 | -34.9788 | -28.6682
18.3128 | -34.9793 | -28.6755
18.3189 | -34.9798 | -28.683
18.325 | -34.9803 | -28.69
18.3311 | -34.9808 | -28.6976
18.3372 | -34.9813 | -28.7047
18.3433 | -34.9817 | -28.7118
18.3494 | -34.9822 | -28.7191
18.3555 | -34.9827 | -28.7263
18.3616 | -34.9832 | -28.7333
18.3677 | -34.9837 | -28.7405
18.3738 | -34.9842 | -28.7478
18.3799 | -34.9846 | -28.7549
18.386 | -34.9851 | -28.7622
18.3921 | -34.9856 | -28.7695
18.3982 | -34.9861 | -28.7768
18.4043 | -34.9865 | -28.784
18.4105 | -34.987 | -28.7915
18.4166 | -34.9875 | -28.7988
18.4227 | -34.9879 | -28.8062
18.4288 | -34.9885 | -28.8138
18.4349 | -34.9889 | -28.821
18.441 | -34.9894 | -28.8283
18.4471 | -34.9899 | -28.8356
18.4532 | -34.9904 | -28.8431
18.4593 | -34.9909 | -28.8503
18.4654 | -34.9913 | -28.8573
18.4715 | -34.9918 | -28.8647
18.4776 | -34.9923 | -28.8723
18.4837 | -34.9928 | -28.8796
18.4898 | -34.9933 | -28.8871
18.4959 | -34.9938 | -28.8945
18.502 | -34.9943 | -28.902
18.5081 | -34.9948 | -28.9097
18.5142 | -34.9953 | -28.9172
18.5203 | -34.9958 | -28.9242
18.5264 | -34.9962 | -28.9314
18.5325 | -34.9968 | -28.939
18.5386 | -34.9973 | -28.9466
18.5447 | -34.9978 | -28.9539
18.5508 | -34.9983 | -28.9612
18.557 | -34.9988 | -28.9688
18.5631 | -34.9993 | -28.9761
18.5692 | -34.9997 | -28.9832
18.5753 | -35.0002 | -28.9903
18.5814 | -35.0007 | -28.9978
18.5875 | -35.0012 | -29.0054
18.5936 | -35.0017 | -29.0131
18.5997 | -35.0022 | -29.0205
18.6058 | -35.0026 | -29.0276
18.6119 | -35.0031 | -29.035
18.618 | -35.0036 | -29.0423
18.6241 | -35.0041 | -29.0498
18.6302 | -35.0046 | -29.0572
18.6363 | -35.0051 | -29.0644
18.6424 | -35.0056 | -29.0715
18.6485 | -35.0061 | -29.0789
18.6546 | -35.0066 | -29.0861
18.6607 | -35.007 | -29.0932
18.6668 | -35.0076 | -29.1006
18.6729 | -35.008 | -29.108
18.679 | -35.0086 | -29.1151
18.6851 | -35.009 | -29.1223
18.6912 | -35.0095 | -29.1295
18.6974 | -35.01 | -29.1367
18.7035 | -35.0105 | -29.1442
18.7096 | -35.011 | -29.1515
18.7157 | -35.0115 | -29.1588
18.7218 | -35.012 | -29.1661
18.7279 | -35.0125 | -29.1732
18.734 | -35.013 | -29.1806
18.7401 | -35.0135 | -29.1876
18.7462 | -35.014 | -29.1945
18.7523 | -35.0146 | -29.2017
18.7584 | -35.0151 | -29.2093
18.7645 | -35.0156 | -29.2164
18.7706 | -35.016 | -29.2237
18.7767 | -35.0166 | -29.2309
18.7828 | -35.017 | -29.2379
18.7889 | -35.0175 | -29.245
18.795 | -35.018 | -29.2522
18.8011 | -35.0185 | -29.2596
18.8072 | -35.019 | -29.267
18.8133 | -35.0195 | -29.2744
18.8194 | -35.02 | -29.2815
18.8255 | -35.0205 | -29.2887
18.8316 | -35.021 | -29.296
18.8377 | -35.0214 | -29.3032
18.8439 | -35.0219 | -29.3103
18.85 | -35.0224 | -29.3177
18.8561 | -35.0229 | -29.3249
18.8622 | -35.0234 | -29.3322
18.8683 | -35.0239 | -29.3393
18.8744 | -35.0244 | -29.3464
18.8805 | -35.0249 | -29.3537
18.8866 | -35.0254 | -29.3611
18.8927 | -35.0259 | -29.3681
18.8988 | -35.0264 | -29.3755
18.9049 | -35.0269 | -29.3827
18.911 | -35.0274 | -29.3899
18.9171 | -35.0279 | -29.3971
18.9232 | -35.0285 | -29.4042
18.9293 | -35.0289 | -29.4115
18.9354 | -35.0294 | -29.4184
18.9415 | -35.0299 | -29.4255
18.9476 | -35.0304 | -29.4328
18.9537 | -35.0309 | -29.4403
18.9598 | -35.0314 | -29.447
18.9659 | -35.032 | -29.4543
18.972 | -35.0324 | -29.4615
18.9781 | -35.0329 | -29.4685
18.9843 | -35.0334 | -29.4757
18.9904 | -35.0338 | -29.4831
18.9965 | -35.0344 | -29.4902
19.0026 | -35.0349 | -29.4975
19.0087 | -35.0354 | -29.5048
19.0148 | -35.0359 | -29.512
19.0209 | -35.0364 | -29.5198
19.027 | -35.0369 | -29.5272
19.0331 | -35.0374 | -29.5345
19.0392 | -35.038 | -29.5417
19.0453 | -35.0385 | -29.549
19.0514 | -35.0389 | -29.5561
19.0575 | -35.0394 | -29.5632
19.0636 | -35.0399 | -29.57
19.0697 | -35.0405 | -29.5776
19.0758 | -35.041 | -29.5846
19.0819 | -35.0415 | -29.592
19.088 | -35.042 | -29.599
19.0941 | -35.0425 | -29.6058
19.1002 | -35.043 | -29.6129
19.1063 | -35.0435 | -29.6202
19.1124 | -35.044 | -29.6273
19.1185 | -35.0445 | -29.6347
19.1246 | -35.0449 | -29.6417
19.1308 | -35.0455 | -29.6486
19.1369 | -35.046 | -29.6556
19.143 | -35.0464 | -29.6628
19.1491 | -35.047 | -29.6699
19.1552 | -35.0474 | -29.6776
19.1613 | -35.0479 | -29.6846
19.1674 | -35.0484 | -29.6919
19.1735 | -35.0489 | -29.6991
19.1796 | -35.0495 | -29.7059
19.1857 | -35.0499 | -29.7132
19.1918 | -35.0504 | -29.7205
19.1979 | -35.0509 | -29.7278
19.204 | -35.0514 | -29.7348
19.2101 | -35.0519 | -29.7421
19.2162 | -35.0524 | -29.7495
19.2223 | -35.0529 | -29.7562
19.2284 | -35.0534 | -29.7638
19.2345 | -35.054 | -29.7711
19.2406 | -35.0545 | -29.778
19.2467 | -35.0549 | -29.7852
19.2528 | -35.0554 | -29.7925
19.2589 | -35.0559 | -29.7997
19.265 | -35.0564 | -29.807
19.2712 | -35.0569 | -29.8142
19.2773 | -35.0574 | -29.8212
19.2834 | -35.0579 | -29.8281
19.2895 | -35.0584 | -29.8354
19.2956 | -35.0589 | -29.8427
19.3017 | -35.0594 | -29.8497
19.3078 | -35.0599 | -29.8569
19.3139 | -35.0604 | -29.8639
19.32 | -35.0609 | -29.8709
19.3261 | -35.0614 | -29.878
19.3322 | -35.062 | -29.8851
19.3383 | -35.0625 | -29.892
19.3444 | -35.063 | -29.899
19.3505 | -35.0635 | -29.906
19.3566 | -35.064 | -29.9131
19.3627 | -35.0645 | -29.9205
19.3688 | -35.065 | -29.9277
19.3749 | -35.0655 | -29.935
19.381 | -35.066 | -29.9417
19.3871 | -35.0665 | -29.9488
19.3932 | -35.067 | -29.9562
19.3993 | -35.0675 | -29.9634
19.4054 | -35.068 | -29.9707
19.4115 | -35.0685 | -29.9779
19.4177 | -35.069 | -29.9849
19.4238 | -35.0695 | -29.9919
19.4299 | -35.07 | -29.9989
19.436 | -35.0705 | -30.0065
19.4421 | -35.0711 | -30.0131
19.4482 | -35.0716 | -30.0204
19.4543 | -35.0721 | -30.0272
19.4604 | -35.0726 | -30.0342
19.4665 | -35.0731 | -30.0412
19.4726 | -35.0736 | -30.0481
19.4787 | -35.0741 | -30.0552
19.4848 | -35.0746 | -30.0623
19.4909 | -35.0751 | -30.0693
19.497 | -35.0756 | -30.0763
19.5031 | -35.0761 | -30.0834
19.5092 | -35.0765 | -30.0905
19.5153 | -35.077 | -30.0976
19.5214 | -35.0775 | -30.1045
19.5275 | -35.078 | -30.1114
19.5336 | -35.0785 | -30.1183
19.5397 | -35.079 | -30.1253
19.5458 | -35.0795 | -30.1325
19.5519 | -35.08 | -30.1394
19.5581 | -35.0805 | -30.1462
19.5642 | -35.081 | -30.1531
19.5703 | -35.0815 | -30.16
19.5764 | -35.082 | -30.1672
19.5825 | -35.0825 | -30.1742
19.5886 | -35.083 | -30.1813
19.5947 | -35.0835 | -30.1882
19.6008 | -35.084 | -30.1953
19.6069 | -35.0845 | -30.2024
19.613 | -35.085 | -30.2096
19.6191 | -35.0855 | -30.2168
19.6252 | -35.086 | -30.2238
19.6313 | -35.0865 | -30.2305
19.6374 | -35.087 | -30.2377
19.6435 | -35.0874 | -30.2445
19.6496 | -35.088 | -30.2514
19.6557 | -35.0884 | -30.2585
19.6618 | -35.0889 | -30.2655
19.6679 | -35.0894 | -30.2726
19.674 | -35.0899 | -30.2794
19.6801 | -35.0904 | -30.2863
19.6862 | -35.091 | -30.2931
19.6923 | -35.0915 | -30.3
19.6984 | -35.092 | -30.3067
19.7046 | -35.0925 | -30.3138
19.7107 | -35.093 | -30.3206
19.7168 | -35.0935 | -30.3279
19.7229 | -35.0939 | -30.335
19.729 | -35.0944 | -30.3423
19.7351 | -35.0949 | -30.3489
19.7412 | -35.0955 | -30.3559
19.7473 | -35.0959 | -30.3629
19.7534 | -35.0964 | -30.3702
19.7595 | -35.0969 | -30.3773
19.7656 | -35.0974 | -30.3848
19.7717 | -35.0979 | -30.3914
19.7778 | -35.0984 | -30.3984
19.7839 | -35.0989 | -30.4053
19.79 | -35.0994 | -30.4123
19.7961 | -35.0998 | -30.4195
19.8022 | -35.1003 | -30.4267
19.8083 | -35.1008 | -30.4338
19.8144 | -35.1014 | -30.4408
19.8205 | -35.1018 | -30.4478
19.8266 | -35.1024 | -30.4547
19.8327 | -35.1028 | -30.4618
19.8388 | -35.1034 | -30.4687
19.845 | -35.1039 | -30.4757
19.8511 | -35.1044 | -30.483
19.8572 | -35.1049 | -30.4902
19.8633 | -35.1053 | -30.4972
19.8694 | -35.1059 | -30.5043
19.8755 | -35.1064 | -30.5113
19.8816 | -35.1069 | -30.5184
19.8877 | -35.1074 | -30.5252
19.8938 | -35.1079 | -30.5323
19.8999 | -35.1084 | -30.5393
19.906 | -35.1089 | -30.5463
19.9121 | -35.1095 | -30.553
19.9182 | -35.11 | -30.5599
19.9243 | -35.1105 | -30.567
19.9304 | -35.1109 | -30.5742
19.9365 | -35.1114 | -30.5814
19.9426 | -35.1119 | -30.5883
19.9487 | -35.1124 | -30.5949
19.9548 | -35.1129 | -30.6021
19.9609 | -35.1135 | -30.609
19.967 | -35.114 | -30.6158
19.9731 | -35.1144 | -30.6228
19.9792 | -35.1149 | -30.6292
19.9853 | -35.1155 | -30.6364
19.9915 | -35.116 | -30.6435
19.9976 | -35.1165 | -30.6505
20.0037 | -35.117 | -30.6574
20.0098 | -35.1175 | -30.6643
20.0159 | -35.118 | -30.6713
20.022 | -35.1185 | -30.6786
20.0281 | -35.119 | -30.6858
20.0342 | -35.1195 | -30.6924
20.0403 | -35.12 | -30.6992
20.0464 | -35.1205 | -30.7065
20.0525 | -35.1211 | -30.7134
20.0586 | -35.1216 | -30.7202
20.0647 | -35.1221 | -30.7275
20.0708 | -35.1226 | -30.735
20.0769 | -35.1231 | -30.7419
20.083 | -35.1236 | -30.7483
20.0891 | -35.124 | -30.7552
20.0952 | -35.1246 | -30.7618
20.1013 | -35.1251 | -30.7689
20.1074 | -35.1256 | -30.7761
20.1135 | -35.1261 | -30.7829
20.1196 | -35.1266 | -30.7901
20.1257 | -35.1271 | -30.7971
20.1319 | -35.1276 | -30.804
20.138 | -35.1281 | -30.8107
20.1441 | -35.1286 | -30.8174
20.1502 | -35.1291 | -30.8247
20.1563 | -35.1296 | -30.8314
20.1624 | -35.1301 | -30.8381
20.1685 | -35.1306 | -30.845
20.1746 | -35.1311 | -30.8523
20.1807 | -35.1316 | -30.8592
20.1868 | -35.1321 | -30.8663
20.1929 | -35.1326 | -30.8732
20.199 | -35.1331 | -30.8801
20.2051 | -35.1336 | -30.8868
20.2112 | -35.1341 | -30.8935
20.2173 | -35.1345 | -30.9007
20.2234 | -35.135 | -30.9076
20.2295 | -35.1355 | -30.9142
20.2356 | -35.136 | -30.9214
20.2417 | -35.1364 | -30.9283
20.2478 | -35.1369 | -30.9353
20.2539 | -35.1374 | -30.9426
20.26 | -35.1378 | -30.9492
20.2661 | -35.1383 | -30.9562
20.2723 | -35.1388 | -30.9631
20.2784 | -35.1393 | -30.97
20.2845 | -35.1398 | -30.977
20.2906 | -35.1403 | -30.9841
20.2967 | -35.1408 | -30.9907
20.3028 | -35.1413 | -30.9975
20.3089 | -35.1418 | -31.0045
20.315 | -35.1423 | -31.0111
20.3211 | -35.1427 | -31.0181
20.3272 | -35.1433 | -31.0248
20.3333 | -35.1438 | -31.0313
20.3394 | -35.1443 | -31.0382
20.3455 | -35.1448 | -31.0452
20.3516 | -35.1453 | -31.0519
20.3577 | -35.1458 | -31.0587
20.3638 | -35.1463 | -31.0654
20.3699 | -35.1468 | -31.0718
20.376 | -35.1473 | -31.0786
20.3821 | -35.1478 | -31.0854
20.3882 | -35.1483 | -31.0926
20.3943 | -35.1488 | -31.0998
20.4004 | -35.1492 | -31.1064
20.4065 | -35.1498 | -31.1131
20.4126 | -35.1503 | -31.1201
20.4188 | -35.1508 | -31.1272
20.4249 | -35.1513 | -31.1343
20.431 | -35.1517 | -31.1412
20.4371 | -35.1522 | -31.1476
20.4432 | -35.1527 | -31.1544
20.4493 | -35.1532 | -31.1612
20.4554 | -35.1537 | -31.1681
20.4615 | -35.1542 | -31.1751
20.4676 | -35.1547 | -31.182
20.4737 | -35.1552 | -31.1885
20.4798 | -35.1558 | -31.1953
20.4859 | -35.1562 | -31.2018
20.492 | -35.1567 | -31.2087
20.4981 | -35.1573 | -31.2155
20.5042 | -35.1578 | -31.2221
20.5103 | -35.1583 | -31.229
20.5164 | -35.1588 | -31.236
20.5225 | -35.1592 | -31.2428
20.5286 | -35.1597 | -31.2499
20.5347 | -35.1602 | -31.2567
20.5408 | -35.1607 | -31.2635
20.5469 | -35.1612 | -31.2703
20.553 | -35.1617 | -31.2771
20.5592 | -35.1622 | -31.284
20.5653 | -35.1626 | -31.2911
20.5714 | -35.1631 | -31.2976
20.5775 | -35.1636 | -31.3042
20.5836 | -35.1641 | -31.3112
20.5897 | -35.1646 | -31.3181
20.5958 | -35.1651 | -31.3249
20.6019 | -35.1656 | -31.3318
20.608 | -35.1661 | -31.3386
20.6141 | -35.1666 | -31.3452
20.6202 | -35.167 | -31.3519
20.6263 | -35.1675 | -31.3585
20.6324 | -35.168 | -31.3651
20.6385 | -35.1685 | -31.3719
20.6446 | -35.169 | -31.3788
20.6507 | -35.1694 | -31.3856
20.6568 | -35.1699 | -31.3922
20.6629 | -35.1704 | -31.3988
20.669 | -35.1709 | -31.4055
20.6751 | -35.1713 | -31.4122
20.6812 | -35.1718 | -31.4189
20.6873 | -35.1723 | -31.4257
20.6934 | -35.1728 | -31.4327
20.6995 | -35.1733 | -31.4393
20.7057 | -35.1738 | -31.446
20.7118 | -35.1743 | -31.4526
20.7179 | -35.1747 | -31.4594
20.724 | -35.1752 | -31.4658
20.7301 | -35.1757 | -31.4726
20.7362 | -35.1762 | -31.4794
20.7423 | -35.1767 | -31.486
20.7484 | -35.1772 | -31.493
20.7545 | -35.1777 | -31.4996
20.7606 | -35.1782 | -31.5069
20.7667 | -35.1787 | -31.5135
20.7728 | -35.1792 | -31.52
20.7789 | -35.1798 | -31.5268
20.785 | -35.1803 | -31.5337
20.7911 | -35.1807 | -31.5405
20.7972 | -35.1812 | -31.5475
20.8033 | -35.1816 | -31.554
20.8094 | -35.1821 | -31.5609
20.8155 | -35.1826 | -31.5676
20.8216 | -35.1831 | -31.5743
20.8277 | -35.1837 | -31.5804
20.8338 | -35.1841 | -31.5874
20.8399 | -35.1846 | -31.5943
20.8461 | -35.1851 | -31.601
20.8522 | -35.1856 | -31.6076
20.8583 | -35.1861 | -31.6139
20.8644 | -35.1866 | -31.621
20.8705 | -35.1872 | -31.6279
20.8766 | -35.1876 | -31.6349
20.8827 | -35.1882 | -31.6417
20.8888 | -35.1886 | -31.6483
20.8949 | -35.1891 | -31.6553
20.901 | -35.1896 | -31.6617
20.9071 | -35.1901 | -31.6682
20.9132 | -35.1906 | -31.675
20.9193 | -35.191 | -31.6818
20.9254 | -35.1915 | -31.6884
20.9315 | -35.192 | -31.6948
20.9376 | -35.1924 | -31.7014
20.9437 | -35.1929 | -31.7083
20.9498 | -35.1934 | -31.7148
20.9559 | -35.194 | -31.7216
20.962 | -35.1945 | -31.7286
20.9681 | -35.195 | -31.7353
20.9742 | -35.1955 | -31.7422
20.9803 | -35.1959 | -31.749
20.9864 | -35.1964 | -31.7559
20.9926 | -35.1968 | -31.7628
20.9987 | -35.1974 | -31.7693
21.0048 | -35.1978 | -31.7763
21.0109 | -35.1983 | -31.7832
21.017 | -35.1988 | -31.7898
21.0231 | -35.1993 | -31.7962
21.0292 | -35.1998 | -31.8029
21.0353 | -35.2003 | -31.8094
21.0414 | -35.2008 | -31.8157
21.0475 | -35.2013 | -31.8225
21.0536 | -35.2018 | -31.8291
21.0597 | -35.2023 | -31.8356
21.0658 | -35.2027 | -31.842
21.0719 | -35.2031 | -31.8482
21.078 | -35.2036 | -31.8553
21.0841 | -35.2041 | -31.8619
21.0902 | -35.2045 | -31.8687
21.0963 | -35.205 | -31.8755
21.1024 | -35.2055 | -31.8824
21.1085 | -35.2058 | -31.889
21.1146 | -35.2063 | -31.8959
21.1207 | -35.2068 | -31.9025
21.1268 | -35.2072 | -31.909
21.133 | -35.2077 | -31.9155
21.1391 | -35.2081 | -31.9219
21.1452 | -35.2086 | -31.9286
21.1513 | -35.2091 | -31.9352
21.1574 | -35.2096 | -31.9421
21.1635 | -35.21 | -31.9487
21.1696 | -35.2105 | -31.9553
21.1757 | -35.2109 | -31.962
21.1818 | -35.2114 | -31.969
21.1879 | -35.2119 | -31.9758
21.194 | -35.2123 | -31.9822
21.2001 | -35.2128 | -31.9888
21.2062 | -35.2133 | -31.9952
21.2123 | -35.2138 | -32.0015
21.2184 | -35.2142 | -32.0082
21.2245 | -35.2147 | -32.0149
21.2306 | -35.2152 | -32.0214
21.2367 | -35.2157 | -32.0279
21.2428 | -35.2161 | -32.0346
21.2489 | -35.2166 | -32.0411
21.255 | -35.2171 | -32.0478
21.2611 | -35.2175 | -32.0544
21.2672 | -35.218 | -32.0614
21.2733 | -35.2184 | -32.0681
21.2795 | -35.2188 | -32.0744
21.2856 | -35.2193 | -32.0811
21.2917 | -35.2198 | -32.0881
21.2978 | -35.2202 | -32.0948
21.3039 | -35.2207 | -32.1011
21.31 | -35.2211 | -32.1078
21.3161 | -35.2216 | -32.1142
21.3222 | -35.2221 | -32.121
21.3283 | -35.2226 | -32.1275
21.3344 | -35.223 | -32.1341
21.3405 | -35.2235 | -32.1411
21.3466 | -35.2241 | -32.1474
21.3527 | -35.2246 | -32.1538
21.3588 | -35.2251 | -32.1599
21.3649 | -35.2256 | -32.1662
21.371 | -35.226 | -32.1724
21.3771 | -35.2265 | -32.1786
21.3832 | -35.2269 | -32.1848
21.3893 | -35.2274 | -32.1906
21.3954 | -35.2278 | -32.1973
21.4015 | -35.2283 | -32.2036
21.4076 | -35.2288 | -32.21
21.4137 | -35.2293 | -32.2165
21.4199 | -35.2298 | -32.2233
21.426 | -35.2302 | -32.2303
21.4321 | -35.2307 | -32.2369
21.4382 | -35.2312 | -32.2433
21.4443 | -35.2316 | -32.2499
21.4504 | -35.2321 | -32.2569
21.4565 | -35.2326 | -32.263
21.4626 | -35.2331 | -32.2695
21.4687 | -35.2336 | -32.2759
21.4748 | -35.2341 | -32.2824
21.4809 | -35.2345 | -32.2887
21.487 | -35.235 | -32.2953
21.4931 | -35.2354 | -32.3018
21.4992 | -35.2359 | -32.3081
21.5053 | -35.2363 | -32.3148
21.5114 | -35.2368 | -32.3213
21.5175 | -35.2372 | -32.3277
21.5236 | -35.2377 | -32.3342
21.5297 | -35.2382 | -32.3404
21.5358 | -35.2387 | -32.3466
21.5419 | -35.2391 | -32.3536
21.548 | -35.2396 | -32.3602
21.5541 | -35.24 | -32.3666
21.5602 | -35.2405 | -32.3731
21.5664 | -35.241 | -32.3797
21.5725 | -35.2414 | -32.3865
21.5786 | -35.2419 | -32.3931
21.5847 | -35.2423 | -32.3997
21.5908 | -35.2428 | -32.4059
21.5969 | -35.2433 | -32.4122
21.603 | -35.2438 | -32.4185
21.6091 | -35.2442 | -32.4249
21.6152 | -35.2447 | -32.4314
21.6213 | -35.2452 | -32.4382
21.6274 | -35.2457 | -32.4447
21.6335 | -35.2462 | -32.4513
21.6396 | -35.2467 | -32.4579
21.6457 | -35.2472 | -32.4648
21.6518 | -35.2476 | -32.4713
21.6579 | -35.248 | -32.4774
21.664 | -35.2485 | -32.484
21.6701 | -35.249 | -32.4906
21.6762 | -35.2495 | -32.4969
21.6823 | -35.2499 | -32.5034
21.6884 | -35.2504 | -32.5098
21.6945 | -35.2508 | -32.5163
21.7006 | -35.2512 | -32.5231
21.7068 | -35.2517 | -32.5294
21.7129 | -35.2522 | -32.5358
21.719 | -35.2527 | -32.5426
21.7251 | -35.2533 | -32.549
21.7312 | -35.2537 | -32.5554
21.7373 | -35.2542 | -32.5614
21.7434 | -35.2547 | -32.5677
21.7495 | -35.2551 | -32.5744
21.7556 | -35.2556 | -32.5813
21.7617 | -35.2561 | -32.5877
21.7678 | -35.2565 | -32.5944
21.7739 | -35.2569 | -32.6006
21.78 | -35.2574 | -32.607
21.7861 | -35.2579 | -32.6132
21.7922 | -35.2584 | -32.6198
21.7983 | -35.2589 | -32.6261
21.8044 | -35.2593 | -32.6327
21.8105 | -35.2598 | -32.6392
21.8166 | -35.2603 | -32.6457
21.8227 | -35.2608 | -32.6522
21.8288 | -35.2612 | -32.6587
21.8349 | -35.2617 | -32.6649
21.841 | -35.2622 | -32.6712
21.8471 | -35.2626 | -32.6775
21.8533 | -35.263 | -32.684
21.8594 | -35.2635 | -32.6907
21.8655 | -35.264 | -32.6971
21.8716 | -35.2645 | -32.7037
21.8777 | -35.265 | -32.7099
21.8838 | -35.2655 | -32.7159
21.8899 | -35.2659 | -32.7226
21.896 | -35.2664 | -32.729
21.9021 | -35.2668 | -32.7358
21.9082 | -35.2672 | -32.7424
21.9143 | -35.2676 | -32.7487
21.9204 | -35.2682 | -32.7551
21.9265 | -35.2686 | -32.7615
21.9326 | -35.2691 | -32.7678
21.9387 | -35.2696 | -32.7745
21.9448 | -35.2701 | -32.7813
21.9509 | -35.2705 | -32.7871
21.957 | -35.271 | -32.7935
21.9631 | -35.2715 | -32.8001
21.9692 | -35.272 | -32.8066
21.9753 | -35.2725 | -32.8131
21.9814 | -35.2729 | -32.8195
21.9875 | -35.2734 | -32.826
21.9937 | -35.2738 | -32.8329
21.9998 | -35.2743 | -32.8396
22.0059 | -35.2748 | -32.8462
22.012 | -35.2753 | -32.8525
22.0181 | -35.2757 | -32.8583
22.0242 | -35.2762 | -32.8646
22.0303 | -35.2766 | -32.8711
22.0364 | -35.2771 | -32.8774
22.0425 | -35.2775 | -32.8839
22.0486 | -35.278 | -32.8901
22.0547 | -35.2784 | -32.8967
22.0608 | -35.2789 | -32.9031
22.0669 | -35.2794 | -32.9092
22.073 | -35.2799 | -32.9155
22.0791 | -35.2804 | -32.922
22.0852 | -35.2809 | -32.9284
22.0913 | -35.2813 | -32.935
22.0974 | -35.2818 | -32.9413
22.1035 | -35.2822 | -32.9476
22.1096 | -35.2826 | -32.9542
22.1157 | -35.2831 | -32.9605
22.1218 | -35.2835 | -32.9669
22.1279 | -35.284 | -32.9728
22.134 | -35.2845 | -32.979
22.1402 | -35.2849 | -32.9855
22.1463 | -35.2854 | -32.9918
22.1524 | -35.2859 | -32.9984
22.1585 | -35.2863 | -33.0049
22.1646 | -35.2867 | -33.0108
22.1707 | -35.2872 | -33.0176
22.1768 | -35.2877 | -33.0238
22.1829 | -35.2881 | -33.0301
22.189 | -35.2885 | -33.0363
22.1951 | -35.289 | -33.0425
22.2012 | -35.2895 | -33.0487
22.2073 | -35.2899 | -33.0556
22.2134 | -35.2904 | -33.0619
22.2195 | -35.2909 | -33.0682
22.2256 | -35.2913 | -33.0749
22.2317 | -35.2918 | -33.0811
22.2378 | -35.2922 | -33.0874
22.2439 | -35.2927 | -33.0939
22.25 | -35.2931 | -33.1004
22.2561 | -35.2936 | -33.1066
22.2622 | -35.294 | -33.1129
22.2683 | -35.2945 | -33.1187
22.2744 | -35.295 | -33.1248
22.2806 | -35.2954 | -33.1313
22.2867 | -35.2959 | -33.1376
22.2928 | -35.2963 | -33.1441
22.2989 | -35.2968 | -33.1506
22.305 | -35.2972 | -33.1569
22.3111 | -35.2977 | -33.1634
22.3172 | -35.2982 | -33.1699
22.3233 | -35.2986 | -33.1762
22.3294 | -35.2991 | -33.1828
22.3355 | -35.2996 | -33.189
22.3416 | -35.3 | -33.1953
22.3477 | -35.3005 | -33.2017
22.3538 | -35.301 | -33.208
22.3599 | -35.3014 | -33.2141
22.366 | -35.3019 | -33.2202
22.3721 | -35.3023 | -33.2259
22.3782 | -35.3028 | -33.2326
22.3843 | -35.3032 | -33.2392
22.3904 | -35.3037 | -33.2456
22.3965 | -35.3042 | -33.2519
22.4026 | -35.3046 | -33.2579
22.4087 | -35.305 | -33.264
22.4148 | -35.3055 | -33.2703
22.4209 | -35.3059 | -33.2766
22.4271 | -35.3063 | -33.2826
22.4332 | -35.3068 | -33.2887
22.4393 | -35.3072 | -33.295
22.4454 | -35.3076 | -33.3008
22.4515 | -35.3081 | -33.3071
22.4576 | -35.3085 | -33.3134
22.4637 | -35.309 | -33.3197
22.4698 | -35.3095 | -33.3264
22.4759 | -35.3099 | -33.3327
22.482 | -35.3104 | -33.339
22.4881 | -35.3108 | -33.3452
22.4942 | -35.3113 | -33.3513
22.5003 | -35.3117 | -33.3574
22.5064 | -35.3122 | -33.3637
22.5125 | -35.3126 | -33.37
22.5186 | -35.3131 | -33.3763
22.5247 | -35.3135 | -33.3825
22.5308 | -35.3139 | -33.389
22.5369 | -35.3144 | -33.3957
22.543 | -35.3148 | -33.402
22.5491 | -35.3153 | -33.4082
22.5552 | -35.3157 | -33.4146
22.5613 | -35.3162 | -33.4211
22.5675 | -35.3166 | -33.4275
22.5736 | -35.3171 | -33.434
22.5797 | -35.3175 | -33.44
22.5858 | -35.3179 | -33.4462
22.5919 | -35.3184 | -33.4524
22.598 | -35.3188 | -33.4588
22.6041 | -35.3193 | -33.465
22.6102 | -35.3197 | -33.4714
22.6163 | -35.3202 | -33.4773
22.6224 | -35.3207 | -33.4836
22.6285 | -35.3211 | -33.4897
22.6346 | -35.3216 | -33.4958
22.6407 | -35.322 | -33.5022
22.6468 | -35.3225 | -33.5082
22.6529 | -35.3229 | -33.5148
22.659 | -35.3234 | -33.5211
22.6651 | -35.3238 | -33.5274
22.6712 | -35.3243 | -33.5336
22.6773 | -35.3247 | -33.5397
22.6834 | -35.3252 | -33.5461
22.6895 | -35.3256 | -33.5519
22.6956 | -35.3261 | -33.558
22.7017 | -35.3266 | -33.5643
22.7079 | -35.3271 | -33.5703
22.714 | -35.3275 | -33.5766
22.7201 | -35.328 | -33.5829
22.7262 | -35.3283 | -33.5892
22.7323 | -35.3288 | -33.5954
22.7384 | -35.3292 | -33.6017
22.7445 | -35.3297 | -33.6081
22.7506 | -35.3302 | -33.6144
22.7567 | -35.3306 | -33.6209
22.7628 | -35.3311 | -33.6268
22.7689 | -35.3315 | -33.633
22.775 | -35.332 | -33.6394
22.7811 | -35.3325 | -33.6455
22.7872 | -35.3329 | -33.6519
22.7933 | -35.3333 | -33.6577
22.7994 | -35.3338 | -33.664
22.8055 | -35.3342 | -33.6703
22.8116 | -35.3346 | -33.6765
22.8177 | -35.3351 | -33.683
22.8238 | -35.3355 | -33.6894
22.8299 | -35.336 | -33.6952
22.836 | -35.3364 | -33.7015
22.8421 | -35.3369 | -33.7075
22.8482 | -35.3373 | -33.7138
22.8544 | -35.3378 | -33.7201
22.8605 | -35.3382 | -33.7264
22.8666 | -35.3387 | -33.7326
22.8727 | -35.3391 | -33.739
22.8788 | -35.3396 | -33.7453
22.8849 | -35.3401 | -33.7515
22.891 | -35.3405 | -33.7578
22.8971 | -35.341 | -33.7641
22.9032 | -35.3414 | -33.7702
22.9093 | -35.3418 | -33.7764
22.9154 | -35.3423 | -33.7828
22.9215 | -35.3427 | -33.7889
22.9276 | -35.3431 | -33.7951
22.9337 | -35.3435 | -33.8009
22.9398 | -35.344 | -33.8071
22.9459 | -35.3444 | -33.8132
22.952 | -35.3449 | -33.8194
22.9581 | -35.3453 | -33.8259
22.9642 | -35.3458 | -33.8322
22.9703 | -35.3462 | -33.838
22.9764 | -35.3467 | -33.8442
22.9825 | -35.3471 | -33.8504
22.9886 | -35.3476 | -33.8565
22.9948 | -35.348 | -33.8628
23.0009 | -35.3484 | -33.8693
23.007 | -35.3489 | -33.8757
23.0131 | -35.3493 | -33.8818
23.0192 | -35.3498 | -33.8877
23.0253 | -35.3502 | -33.8941
23.0314 | -35.3507 | -33.9002
23.0375 | -35.3512 | -33.9066
23.0436 | -35.3516 | -33.9126
23.0497 | -35.3521 | -33.9186
23.0558 | -35.3525 | -33.9247
23.0619 | -35.353 | -33.9308
23.068 | -35.3534 | -33.9371
23.0741 | -35.3539 | -33.9429
23.0802 | -35.3543 | -33.9489
23.0863 | -35.3547 | -33.955
23.0924 | -35.3552 | -33.9615
23.0985 | -35.3556 | -33.9678
23.1046 | -35.3561 | -33.9739
23.1107 | -35.3565 | -33.98
23.1168 | -35.3569 | -33.9864
23.1229 | -35.3573 | -33.9926
23.129 | -35.3578 | -33.9989
23.1351 | -35.3582 | -34.0052
23.1413 | -35.3587 | -34.0113
23.1474 | -35.3591 | -34.0172
23.1535 | -35.3595 | -34.0232
23.1596 | -35.3599 | -34.0292
23.1657 | -35.3604 | -34.0352
23.1718 | -35.3608 | -34.0412
23.1779 | -35.3612 | -34.0474
23.184 | -35.3617 | -34.053
23.1901 | -35.3622 | -34.0594
23.1962 | -35.3626 | -34.0656
23.2023 | -35.3631 | -34.0717
23.2084 | -35.3636 | -34.078
23.2145 | -35.3641 | -34.084
23.2206 | -35.3645 | -34.0902
23.2267 | -35.3649 | -34.0965
23.2328 | -35.3653 | -34.1027
23.2389 | -35.3658 | -34.1091
23.245 | -35.3663 | -34.1153
23.2511 | -35.3667 | -34.1213
23.2572 | -35.3671 | -34.1275
23.2633 | -35.3676 | -34.1336
23.2694 | -35.368 | -34.1397
23.2755 | -35.3685 | -34.1461
23.2817 | -35.3689 | -34.1521
23.2878 | -35.3693 | -34.1581
23.2939 | -35.3698 | -34.1642
23.3 | -35.3702 | -34.1704
23.3061 | -35.3707 | -34.1767
23.3122 | -35.3712 | -34.1826
23.3183 | -35.3716 | -34.1887
23.3244 | -35.372 | -34.1946
23.3305 | -35.3724 | -34.201
23.3366 | -35.3729 | -34.2073
23.3427 | -35.3733 | -34.213
23.3488 | -35.3738 | -34.219
23.3549 | -35.3742 | -34.2251
23.361 | -35.3747 | -34.2313
23.3671 | -35.3752 | -34.2375
23.3732 | -35.3756 | -34.2433
23.3793 | -35.3761 | -34.2493
23.3854 | -35.3765 | -34.2556
23.3915 | -35.377 | -34.2615
23.3976 | -35.3775 | -34.2678
23.4037 | -35.3779 | -34.274
23.4098 | -35.3783 | -34.28
23.4159 | -35.3787 | -34.2863
23.422 | -35.3792 | -34.2924
23.4282 | -35.3796 | -34.2985
23.4343 | -35.3801 | -34.3044
23.4404 | -35.3805 | -34.3107
23.4465 | -35.3809 | -34.3168
23.4526 | -35.3814 | -34.3232
23.4587 | -35.3818 | -34.3293
23.4648 | -35.3822 | -34.3354
23.4709 | -35.3826 | -34.3412
23.477 | -35.383 | -34.3475
23.4831 | -35.3834 | -34.3537
23.4892 | -35.3838 | -34.36
23.4953 | -35.3843 | -34.366
23.5014 | -35.3847 | -34.3718
23.5075 | -35.3851 | -34.3778
23.5136 | -35.3856 | -34.3837
23.5197 | -35.386 | -34.3897
23.5258 | -35.3864 | -34.3959
23.5319 | -35.3869 | -34.4018
23.538 | -35.3874 | -34.408
23.5441 | -35.3877 | -34.4141
23.5502 | -35.3881 | -34.4203
23.5563 | -35.3886 | -34.4264
23.5624 | -35.389 | -34.4329
23.5686 | -35.3895 | -34.439
23.5747 | -35.3899 | -34.4449
23.5808 | -35.3904 | -34.4509
23.5869 | -35.3908 | -34.4571
23.593 | -35.3913 | -34.4632
23.5991 | -35.3917 | -34.4693
23.6052 | -35.3922 | -34.4751
23.6113 | -35.3926 | -34.4812
23.6174 | -35.3931 | -34.4873
23.6235 | -35.3935 | -34.4933
23.6296 | -35.3939 | -34.4997
23.6357 | -35.3943 | -34.5061
23.6418 | -35.3947 | -34.512
23.6479 | -35.3951 | -34.5181
23.654 | -35.3956 | -34.5243
23.6601 | -35.396 | -34.5303
23.6662 | -35.3964 | -34.5366
23.6723 | -35.3969 | -34.5424
23.6784 | -35.3973 | -34.5487
23.6845 | -35.3978 | -34.5548
23.6906 | -35.3982 | -34.561
23.6967 | -35.3986 | -34.5671
23.7028 | -35.399 | -34.5733
23.7089 | -35.3994 | -34.5795
23.7151 | -35.3998 | -34.5856
23.7212 | -35.4003 | -34.5918
23.7273 | -35.4007 | -34.5979
23.7334 | -35.4012 | -34.604
23.7395 | -35.4016 | -34.6097
23.7456 | -35.402 | -34.6156
23.7517 | -35.4025 | -34.6217
23.7578 | -35.4029 | -34.6277
23.7639 | -35.4033 | -34.6339
23.77 | -35.4038 | -34.64
23.7761 | -35.4042 | -34.6458
23.7822 | -35.4047 | -34.6521
23.7883 | -35.4051 | -34.6581
23.7944 | -35.4056 | -34.6643
23.8005 | -35.406 | -34.6704
23.8066 | -35.4065 | -34.6765
23.8127 | -35.4069 | -34.6827
23.8188 | -35.4074 | -34.6888
23.8249 | -35.4078 | -34.6952
23.831 | -35.4082 | -34.7015
23.8371 | -35.4087 | -34.7074
23.8432 | -35.4091 | -34.7133
23.8493 | -35.4096 | -34.7191
23.8555 | -35.41 | -34.7252
23.8616 | -35.4105 | -34.7313
23.8677 | -35.411 | -34.7375
23.8738 | -35.4114 | -34.7435
23.8799 | -35.4118 | -34.7496
23.886 | -35.4122 | -34.7552
23.8921 | -35.4126 | -34.7614
23.8982 | -35.4131 | -34.7674
23.9043 | -35.4135 | -34.7738
23.9104 | -35.4139 | -34.7799
23.9165 | -35.4143 | -34.7858
23.9226 | -35.4147 | -34.7915
23.9287 | -35.4152 | -34.7976
23.9348 | -35.4156 | -34.8036
23.9409 | -35.416 | -34.8097
23.947 | -35.4164 | -34.8157
23.9531 | -35.4168 | -34.8213
23.9592 | -35.4173 | -34.8273
23.9653 | -35.4177 | -34.8335
23.9714 | -35.4181 | -34.8394
23.9775 | -35.4185 | -34.8455
23.9836 | -35.419 | -34.8516
23.9897 | -35.4194 | -34.8575
23.9958 | -35.4199 | -34.8636
24.002 | -35.4203 | -34.8696
24.0081 | -35.4207 | -34.8755
24.0142 | -35.4212 | -34.8816
24.0203 | -35.4216 | -34.8875
24.0264 | -35.422 | -34.8935
24.0325 | -35.4224 | -34.8995
24.0386 | -35.4228 | -34.9055
24.0447 | -35.4232 | -34.9114
24.0508 | -35.4237 | -34.9172
24.0569 | -35.4242 | -34.9231
24.063 | -35.4246 | -34.929
24.0691 | -35.4251 | -34.9348
24.0752 | -35.4255 | -34.9408
24.0813 | -35.4259 | -34.9466
24.0874 | -35.4264 | -34.9525
24.0935 | -35.4268 | -34.9585
24.0996 | -35.4272 | -34.9642
24.1057 | -35.4277 | -34.9699
24.1118 | -35.4281 | -34.9761
24.1179 | -35.4285 | -34.9822
24.124 | -35.4289 | -34.9879
24.1301 | -35.4294 | -34.9939
24.1362 | -35.4298 | -34.9998
24.1424 | -35.4302 | -35.0058
24.1485 | -35.4307 | -35.0117
24.1546 | -35.4311 | -35.0179
24.1607 | -35.4315 | -35.0238
24.1668 | -35.4319 | -35.0295
24.1729 | -35.4323 | -35.0355
24.179 | -35.4327 | -35.0416
24.1851 | -35.4332 | -35.0476
24.1912 | -35.4336 | -35.0536
24.1973 | -35.4341 | -35.0599
24.2034 | -35.4345 | -35.0656
24.2095 | -35.435 | -35.0718
24.2156 | -35.4354 | -35.078
24.2217 | -35.4358 | -35.0841
24.2278 | -35.4362 | -35.0902
24.2339 | -35.4367 | -35.0961
24.24 | -35.4371 | -35.102
24.2461 | -35.4375 | -35.1081
24.2522 | -35.4379 | -35.114
24.2583 | -35.4383 | -35.1202
24.2644 | -35.4388 | -35.1264
24.2705 | -35.4392 | -35.1324
24.2766 | -35.4397 | -35.1383
24.2827 | -35.4401 | -35.1439
24.2889 | -35.4405 | -35.1498
24.295 | -35.441 | -35.1559
24.3011 | -35.4414 | -35.1618
24.3072 | -35.4418 | -35.1674
24.3133 | -35.4423 | -35.1732
24.3194 | -35.4426 | -35.1795
24.3255 | -35.4431 | -35.1856
24.3316 | -35.4435 | -35.1914
24.3377 | -35.4439 | -35.1974
24.3438 | -35.4443 | -35.2035
24.3499 | -35.4447 | -35.2095
24.356 | -35.4451 | -35.2155
24.3621 | -35.4455 | -35.2213
24.3682 | -35.446 | -35.227
24.3743 | -35.4464 | -35.2331
24.3804 | -35.4468 | -35.2386
24.3865 | -35.4472 | -35.2446
24.3926 | -35.4476 | -35.2506
24.3987 | -35.448 | -35.2567
24.4048 | -35.4484 | -35.2625
24.4109 | -35.4488 | -35.2684
24.417 | -35.4493 | -35.2743
24.4231 | -35.4497 | -35.2802
24.4293 | -35.4501 | -35.2859
24.4354 | -35.4506 | -35.2915
24.4415 | -35.451 | -35.2974
24.4476 | -35.4514 | -35.3032
24.4537 | -35.4519 | -35.309
24.4598 | -35.4523 | -35.3152
24.4659 | -35.4528 | -35.3213
24.472 | -35.4531 | -35.3273
24.4781 | -35.4536 | -35.3334
24.4842 | -35.454 | -35.339
24.4903 | -35.4544 | -35.3449
24.4964 | -35.4548 | -35.3508
24.5025 | -35.4552 | -35.357
24.5086 | -35.4557 | -35.3629
24.5147 | -35.4561 | -35.3687
24.5208 | -35.4566 | -35.3745
24.5269 | -35.457 | -35.3802
24.533 | -35.4574 | -35.3861
24.5391 | -35.4578 | -35.392
24.5452 | -35.4583 | -35.3979
24.5513 | -35.4587 | -35.4038
24.5574 | -35.4591 | -35.4099
24.5635 | -35.4595 | -35.4154
24.5696 | -35.4599 | -35.4213
24.5758 | -35.4604 | -35.4273
24.5819 | -35.4608 | -35.4332
24.588 | -35.4612 | -35.4389
24.5941 | -35.4617 | -35.4447
24.6002 | -35.4621 | -35.4505
24.6063 | -35.4625 | -35.4564
24.6124 | -35.463 | -35.4625
24.6185 | -35.4633 | -35.4685
24.6246 | -35.4638 | -35.4744
24.6307 | -35.4642 | -35.4806
24.6368 | -35.4646 | -35.4864
24.6429 | -35.4651 | -35.4923
24.649 | -35.4655 | -35.4983
24.6551 | -35.4659 | -35.5041
24.6612 | -35.4664 | -35.5096
24.6673 | -35.4668 | -35.5157
24.6734 | -35.4672 | -35.5215
24.6795 | -35.4676 | -35.5275
24.6856 | -35.468 | -35.5336
24.6917 | -35.4685 | -35.5394
24.6978 | -35.4689 | -35.5453
24.7039 | -35.4693 | -35.5511
24.71 | -35.4697 | -35.5572
24.7162 | -35.4702 | -35.5631
24.7223 | -35.4706 | -35.5689
24.7284 | -35.471 | -35.5745
24.7345 | -35.4714 | -35.5804
24.7406 | -35.4719 | -35.5863
24.7467 | -35.4723 | -35.5922
24.7528 | -35.4727 | -35.5981
24.7589 | -35.4731 | -35.6039
24.765 | -35.4735 | -35.6095
24.7711 | -35.4739 | -35.6152
24.7772 | -35.4743 | -35.6211
24.7833 | -35.4747 | -35.6271
24.7894 | -35.4752 | -35.6329
24.7955 | -35.4756 | -35.6389
24.8016 | -35.476 | -35.6444
24.8077 | -35.4764 | -35.6503
24.8138 | -35.4768 | -35.6561
24.8199 | -35.4772 | -35.6619
24.826 | -35.4777 | -35.6677
24.8321 | -35.4781 | -35.6736
24.8382 | -35.4785 | -35.6792
24.8443 | -35.4789 | -35.6851
24.8504 | -35.4793 | -35.6907
24.8565 | -35.4798 | -35.6966
24.8627 | -35.4802 | -35.7025
24.8688 | -35.4806 | -35.7084
24.8749 | -35.481 | -35.7143
24.881 | -35.4814 | -35.72
24.8871 | -35.4819 | -35.7259
24.8932 | -35.4823 | -35.7316
24.8993 | -35.4828 | -35.7373
24.9054 | -35.4832 | -35.743
24.9115 | -35.4837 | -35.7491
24.9176 | -35.4841 | -35.7551
24.9237 | -35.4845 | -35.7608
24.9298 | -35.4849 | -35.7666
24.9359 | -35.4854 | -35.7725
24.942 | -35.4858 | -35.778
24.9481 | -35.4862 | -35.784
24.9542 | -35.4866 | -35.7901
24.9603 | -35.487 | -35.7961
24.9664 | -35.4874 | -35.8019
24.9725 | -35.4878 | -35.8078
24.9786 | -35.4882 | -35.8131
24.9847 | -35.4886 | -35.8187
24.9908 | -35.489 | -35.8245
24.9969 | -35.4894 | -35.8306
25.0031 | -35.4898 | -35.8365
25.0092 | -35.4902 | -35.842
25.0153 | -35.4906 | -35.8476
25.0214 | -35.491 | -35.8536
25.0275 | -35.4914 | -35.8592
25.0336 | -35.4919 | -35.8651
25.0397 | -35.4923 | -35.8711
25.0458 | -35.4927 | -35.8767
25.0519 | -35.4931 | -35.8827
25.058 | -35.4935 | -35.8885
25.0641 | -35.4939 | -35.8944
25.0702 | -35.4944 | -35.9001
25.0763 | -35.4948 | -35.9058
25.0824 | -35.4952 | -35.9114
25.0885 | -35.4956 | -35.9173
25.0946 | -35.496 | -35.9231
25.1007 | -35.4964 | -35.9289
25.1068 | -35.4969 | -35.9346
25.1129 | -35.4973 | -35.9403
25.119 | -35.4976 | -35.9458
25.1251 | -35.4981 | -35.9517
25.1312 | -35.4984 | -35.9577
25.1373 | -35.4989 | -35.9636
25.1435 | -35.4993 | -35.9696
25.1496 | -35.4997 | -35.9752
25.1557 | -35.5001 | -35.9809
25.1618 | -35.5006 | -35.9869
25.1679 | -35.5009 | -35.9926
25.174 | -35.5013 | -35.9985
25.1801 | -35.5018 | -36.0044
25.1862 | -35.5021 | -36.0102
25.1923 | -35.5025 | -36.0159
25.1984 | -35.503 | -36.0215
25.2045 | -35.5034 | -36.0275
25.2106 | -35.5038 | -36.0333
25.2167 | -35.5042 | -36.039
25.2228 | -35.5046 | -36.0445
25.2289 | -35.505 | -36.0503
25.235 | -35.5055 | -36.056
25.2411 | -35.5058 | -36.0618
25.2472 | -35.5062 | -36.0675
25.2533 | -35.5066 | -36.0734
25.2594 | -35.507 | -36.0787
25.2655 | -35.5074 | -36.0844
25.2716 | -35.5078 | -36.0903
25.2777 | -35.5083 | -36.0964
25.2838 | -35.5086 | -36.1023
25.29 | -35.5091 | -36.108
25.2961 | -35.5095 | -36.1137
25.3022 | -35.5099 | -36.1196
25.3083 | -35.5103 | -36.1253
25.3144 | -35.5107 | -36.1309
25.3205 | -35.511 | -36.1366
25.3266 | -35.5114 | -36.1422
25.3327 | -35.5118 | -36.1478
25.3388 | -35.5122 | -36.1535
25.3449 | -35.5126 | -36.1594
25.351 | -35.513 | -36.1651
25.3571 | -35.5135 | -36.171
25.3632 | -35.5139 | -36.1767
25.3693 | -35.5143 | -36.1826
25.3754 | -35.5146 | -36.1884
25.3815 | -35.5151 | -36.1939
25.3876 | -35.5155 | -36.1997
25.3937 | -35.5159 | -36.2053
25.3998 | -35.5163 | -36.2107
25.4059 | -35.5167 | -36.2166
25.412 | -35.5171 | -36.2222
25.4181 | -35.5175 | -36.2281
25.4242 | -35.5178 | -36.234
25.4304 | -35.5183 | -36.2397
25.4365 | -35.5187 | -36.2454
25.4426 | -35.5191 | -36.2511
25.4487 | -35.5195 | -36.257
25.4548 | -35.5199 | -36.2628
25.4609 | -35.5203 | -36.2684
25.467 | -35.5207 | -36.2738
25.4731 | -35.5211 | -36.2795
25.4792 | -35.5215 | -36.2855
25.4853 | -35.5219 | -36.2912
25.4914 | -35.5223 | -36.297
25.4975 | -35.5227 | -36.3027
25.5036 | -35.5231 | -36.3082
25.5097 | -35.5235 | -36.3139
25.5158 | -35.524 | -36.3195
25.5219 | -35.5244 | -36.3251
25.528 | -35.5248 | -36.3308
25.5341 | -35.5252 | -36.3365
25.5402 | -35.5255 | -36.3422
25.5463 | -35.5259 | -36.3482
25.5524 | -35.5263 | -36.3542
25.5585 | -35.5268 | -36.3597
25.5646 | -35.5272 | -36.3655
25.5707 | -35.5275 | -36.3712
25.5769 | -35.5279 | -36.377
25.583 | -35.5283 | -36.3828
25.5891 | -35.5287 | -36.3886
25.5952 | -35.5291 | -36.3943
25.6013 | -35.5295 | -36.3999
25.6074 | -35.5299 | -36.4054
25.6135 | -35.5304 | -36.411
25.6196 | -35.5308 | -36.4169
25.6257 | -35.5311 | -36.4223
25.6318 | -35.5315 | -36.4279
25.6379 | -35.5319 | -36.4336
25.644 | -35.5323 | -36.4392
25.6501 | -35.5327 | -36.4447
25.6562 | -35.5331 | -36.4504
25.6623 | -35.5335 | -36.4562
25.6684 | -35.5339 | -36.462
25.6745 | -35.5343 | -36.4677
25.6806 | -35.5347 | -36.4732
25.6867 | -35.5351 | -36.4792
25.6928 | -35.5355 | -36.4851
25.6989 | -35.5358 | -36.4907
25.705 | -35.5362 | -36.4967
25.7111 | -35.5366 | -36.5025
25.7173 | -35.537 | -36.508
25.7234 | -35.5374 | -36.5137
25.7295 | -35.5377 | -36.5195
25.7356 | -35.5381 | -36.5254
25.7417 | -35.5385 | -36.531
25.7478 | -35.5389 | -36.5366
25.7539 | -35.5392 | -36.5423
25.76 | -35.5396 | -36.5479
25.7661 | -35.54 | -36.5535
25.7722 | -35.5404 | -36.5592
25.7783 | -35.5408 | -36.5649
25.7844 | -35.5412 | -36.5704
25.7905 | -35.5416 | -36.5762
25.7966 | -35.542 | -36.5815
25.8027 | -35.5423 | -36.5873
25.8088 | -35.5427 | -36.5932
25.8149 | -35.5431 | -36.5989
25.821 | -35.5434 | -36.6042
25.8271 | -35.5438 | -36.6102
25.8332 | -35.5442 | -36.6159
25.8393 | -35.5446 | -36.6214
25.8454 | -35.545 | -36.6269
25.8515 | -35.5454 | -36.6326
25.8576 | -35.5458 | -36.6382
25.8638 | -35.5462 | -36.644
25.8699 | -35.5466 | -36.6498
25.876 | -35.5469 | -36.6555
25.8821 | -35.5473 | -36.6612
25.8882 | -35.5477 | -36.6666
25.8943 | -35.5481 | -36.6724
25.9004 | -35.5485 | -36.6782
25.9065 | -35.5489 | -36.6839
25.9126 | -35.5493 | -36.6894
25.9187 | -35.5497 | -36.695
25.9248 | -35.5501 | -36.7006
25.9309 | -35.5504 | -36.7063
25.937 | -35.5509 | -36.7121
25.9431 | -35.5513 | -36.7177
25.9492 | -35.5517 | -36.7233
25.9553 | -35.552 | -36.729
25.9614 | -35.5525 | -36.7343
25.9675 | -35.5529 | -36.7401
25.9736 | -35.5533 | -36.7458
25.9797 | -35.5537 | -36.7516
25.9858 | -35.5541 | -36.7572
25.9919 | -35.5544 | -36.7629
25.998 | -35.5548 | -36.7685
26.0042 | -35.5552 | -36.7743
26.0103 | -35.5555 | -36.78
26.0164 | -35.5559 | -36.7856
26.0225 | -35.5563 | -36.7911
26.0286 | -35.5567 | -36.797
26.0347 | -35.557 | -36.8026
26.0408 | -35.5574 | -36.8083
26.0469 | -35.5578 | -36.8139
26.053 | -35.5582 | -36.8196
26.0591 | -35.5586 | -36.8254
26.0652 | -35.559 | -36.8311
26.0713 | -35.5593 | -36.8372
26.0774 | -35.5597 | -36.8427
26.0835 | -35.5601 | -36.8484
26.0896 | -35.5605 | -36.8539
26.0957 | -35.5609 | -36.8595
26.1018 | -35.5612 | -36.865
26.1079 | -35.5616 | -36.8705
26.114 | -35.5619 | -36.8763
26.1201 | -35.5623 | -36.8819
26.1262 | -35.5627 | -36.8876
26.1323 | -35.563 | -36.8931
26.1384 | -35.5634 | -36.899
26.1445 | -35.5638 | -36.9049
26.1507 | -35.5642 | -36.9107
26.1568 | -35.5646 | -36.9163
26.1629 | -35.565 | -36.9218
26.169 | -35.5653 | -36.9273
26.1751 | -35.5657 | -36.9326
26.1812 | -35.5661 | -36.9381
26.1873 | -35.5664 | -36.9437
26.1934 | -35.5669 | -36.9491
26.1995 | -35.5672 | -36.9549
26.2056 | -35.5676 | -36.9604
26.2117 | -35.568 | -36.9662
26.2178 | -35.5684 | -36.9719
26.2239 | -35.5688 | -36.9775
26.23 | -35.5691 | -36.9833
26.2361 | -35.5695 | -36.989
26.2422 | -35.5698 | -36.9946
26.2483 | -35.5702 | -37
26.2544 | -35.5706 | -37.0055
26.2605 | -35.571 | -37.0112
26.2666 | -35.5713 | -37.0167
26.2727 | -35.5717 | -37.0223
26.2788 | -35.572 | -37.0279
26.2849 | -35.5724 | -37.0336
26.2911 | -35.5728 | -37.039
26.2972 | -35.5731 | -37.0446
26.3033 | -35.5735 | -37.0503
26.3094 | -35.5739 | -37.056
26.3155 | -35.5743 | -37.0615
26.3216 | -35.5746 | -37.0672
26.3277 | -35.575 | -37.0725
26.3338 | -35.5754 | -37.078
26.3399 | -35.5757 | -37.0833
26.346 | -35.5761 | -37.0888
26.3521 | -35.5765 | -37.0944
26.3582 | -35.5769 | -37.1
26.3643 | -35.5773 | -37.1055
26.3704 | -35.5777 | -37.1112
26.3765 | -35.5781 | -37.1169
26.3826 | -35.5785 | -37.1224
26.3887 | -35.5789 | -37.128
26.3948 | -35.5792 | -37.1337
26.4009 | -35.5796 | -37.1393
26.407 | -35.58 | -37.145
26.4131 | -35.5804 | -37.1504
26.4192 | -35.5808 | -37.156
26.4253 | -35.5812 | -37.1615
26.4314 | -35.5815 | -37.1669
26.4376 | -35.582 | -37.1728
26.4437 | -35.5823 | -37.1783
26.4498 | -35.5827 | -37.184
26.4559 | -35.5831 | -37.1895
26.462 | -35.5834 | -37.1954
26.4681 | -35.5838 | -37.201
26.4742 | -35.5842 | -37.2068
26.4803 | -35.5846 | -37.2125
26.4864 | -35.5849 | -37.2179
26.4925 | -35.5853 | -37.2233
26.4986 | -35.5857 | -37.2291
26.5047 | -35.5861 | -37.2349
26.5108 | -35.5865 | -37.2404
26.5169 | -35.5868 | -37.2459
26.523 | -35.5872 | -37.2515
26.5291 | -35.5875 | -37.2571
26.5352 | -35.5879 | -37.2627
26.5413 | -35.5882 | -37.2686
26.5474 | -35.5886 | -37.2744
26.5535 | -35.589 | -37.2802
26.5596 | -35.5893 | -37.2856
26.5657 | -35.5897 | -37.2911
26.5718 | -35.5901 | -37.2965
26.578 | -35.5904 | -37.3022
26.5841 | -35.5908 | -37.3078
26.5902 | -35.5912 | -37.3131
26.5963 | -35.5915 | -37.3185
26.6024 | -35.5919 | -37.3241
26.6085 | -35.5922 | -37.3296
26.6146 | -35.5926 | -37.3351
26.6207 | -35.593 | -37.3407
26.6268 | -35.5934 | -37.3464
26.6329 | -35.5937 | -37.3522
26.639 | -35.594 | -37.3579
26.6451 | -35.5943 | -37.3634
26.6512 | -35.5947 | -37.3688
26.6573 | -35.5951 | -37.3744
26.6634 | -35.5954 | -37.3798
26.6695 | -35.5958 | -37.3855
26.6756 | -35.5961 | -37.3912
26.6817 | -35.5965 | -37.3968
26.6878 | -35.5969 | -37.4025
26.6939 | -35.5972 | -37.408
26.7 | -35.5975 | -37.4134
26.7061 | -35.5979 | -37.4188
26.7122 | -35.5982 | -37.4243
26.7183 | -35.5985 | -37.4299
26.7245 | -35.5989 | -37.4355
26.7306 | -35.5992 | -37.441
26.7367 | -35.5996 | -37.4464
26.7428 | -35.6 | -37.4518
26.7489 | -35.6003 | -37.4574
26.755 | -35.6007 | -37.463
26.7611 | -35.6011 | -37.4687
26.7672 | -35.6014 | -37.4742
26.7733 | -35.6018 | -37.4799
26.7794 | -35.6022 | -37.4855
26.7855 | -35.6026 | -37.4911
26.7916 | -35.6029 | -37.4965
26.7977 | -35.6033 | -37.502
26.8038 | -35.6037 | -37.5077
26.8099 | -35.604 | -37.5133
26.816 | -35.6044 | -37.5189
26.8221 | -35.6047 | -37.5243
26.8282 | -35.6051 | -37.5296
26.8343 | -35.6054 | -37.5353
26.8404 | -35.6058 | -37.5409
26.8465 | -35.6061 | -37.5466
26.8526 | -35.6065 | -37.5521
26.8587 | -35.6068 | -37.5576
26.8649 | -35.6072 | -37.5632
26.871 | -35.6075 | -37.5688
26.8771 | -35.6078 | -37.5746
26.8832 | -35.6082 | -37.58
26.8893 | -35.6085 | -37.5857
26.8954 | -35.6089 | -37.5911
26.9015 | -35.6092 | -37.597
26.9076 | -35.6096 | -37.6026
26.9137 | -35.61 | -37.6084
26.9198 | -35.6103 | -37.6142
26.9259 | -35.6107 | -37.6197
26.932 | -35.611 | -37.6254
26.9381 | -35.6114 | -37.6312
26.9442 | -35.6116 | -37.6364
26.9503 | -35.6119 | -37.642
26.9564 | -35.6124 | -37.6476
26.9625 | -35.6127 | -37.6531
26.9686 | -35.6131 | -37.6588
26.9747 | -35.6135 | -37.6643
26.9808 | -35.6138 | -37.6694
26.9869 | -35.6142 | -37.675
26.993 | -35.6145 | -37.6805
26.9991 | -35.6149 | -37.686
27.0052 | -35.6152 | -37.6918
27.0114 | -35.6156 | -37.6976
27.0175 | -35.6159 | -37.7032
27.0236 | -35.6163 | -37.7085
27.0297 | -35.6166 | -37.7141
27.0358 | -35.6169 | -37.7196
27.0419 | -35.6173 | -37.7251
27.048 | -35.6176 | -37.7307
27.0541 | -35.6179 | -37.7364
27.0602 | -35.6183 | -37.7419
27.0663 | -35.6186 | -37.7478
27.0724 | -35.619 | -37.7532
27.0785 | -35.6193 | -37.7586
27.0846 | -35.6197 | -37.7639
27.0907 | -35.62 | -37.7695
27.0968 | -35.6204 | -37.7749
27.1029 | -35.6207 | -37.7803
27.109 | -35.6211 | -37.7858
27.1151 | -35.6215 | -37.7912
27.1212 | -35.6218 | -37.7967
27.1273 | -35.6221 | -37.8025
27.1334 | -35.6225 | -37.808
27.1395 | -35.6228 | -37.8135
27.1456 | -35.6232 | -37.8193
27.1518 | -35.6235 | -37.8247
27.1579 | -35.6239 | -37.8304
27.164 | -35.6242 | -37.8358
27.1701 | -35.6246 | -37.8413
27.1762 | -35.6249 | -37.8467
27.1823 | -35.6252 | -37.8523
27.1884 | -35.6256 | -37.8577
27.1945 | -35.6258 | -37.8632
27.2006 | -35.6262 | -37.8688
27.2067 | -35.6265 | -37.8742
27.2128 | -35.6269 | -37.8797
27.2189 | -35.6272 | -37.8854
27.225 | -35.6275 | -37.8909
27.2311 | -35.6279 | -37.8963
27.2372 | -35.6283 | -37.9018
27.2433 | -35.6287 | -37.9072
27.2494 | -35.629 | -37.9127
27.2555 | -35.6294 | -37.9182
27.2616 | -35.6297 | -37.9238
27.2677 | -35.63 | -37.9292
27.2738 | -35.6304 | -37.9346
27.2799 | -35.6308 | -37.94
27.286 | -35.6311 | -37.9455
27.2921 | -35.6315 | -37.9511
27.2983 | -35.6318 | -37.9566
27.3044 | -35.6321 | -37.9621
27.3105 | -35.6325 | -37.9676
27.3166 | -35.6328 | -37.9729
27.3227 | -35.6331 | -37.9783
27.3288 | -35.6335 | -37.984
27.3349 | -35.6338 | -37.9895
27.341 | -35.6342 | -37.9953
27.3471 | -35.6345 | -38.0005
27.3532 | -35.6348 | -38.0059
27.3593 | -35.6352 | -38.0113
27.3654 | -35.6355 | -38.0169
27.3715 | -35.6359 | -38.022
27.3776 | -35.6362 | -38.0274
27.3837 | -35.6366 | -38.0329
27.3898 | -35.6369 | -38.0384
27.3959 | -35.6373 | -38.044
27.402 | -35.6376 | -38.0493
27.4081 | -35.6379 | -38.0548
27.4142 | -35.6383 | -38.0604
27.4203 | -35.6386 | -38.0659
27.4264 | -35.6389 | -38.0714
27.4325 | -35.6393 | -38.0769
27.4387 | -35.6396 | -38.0825
27.4448 | -35.6399 | -38.088
27.4509 | -35.6403 | -38.0936
27.457 | -35.6406 | -38.0992
27.4631 | -35.6409 | -38.105
27.4692 | -35.6412 | -38.1104
27.4753 | -35.6415 | -38.1156
27.4814 | -35.6418 | -38.1209
27.4875 | -35.6421 | -38.1266
27.4936 | -35.6425 | -38.1319
27.4997 | -35.6428 | -38.1375
27.5058 | -35.6431 | -38.1428
27.5119 | -35.6434 | -38.148
27.518 | -35.6438 | -38.1536
27.5241 | -35.6441 | -38.159
27.5302 | -35.6445 | -38.1645
27.5363 | -35.6448 | -38.1702
27.5424 | -35.6451 | -38.1755
27.5485 | -35.6454 | -38.181
27.5546 | -35.6458 | -38.1864
27.5607 | -35.6462 | -38.1917
27.5668 | -35.6465 | -38.1971
27.5729 | -35.6468 | -38.2027
27.5791 | -35.6471 | -38.2078
27.5852 | -35.6475 | -38.2132
27.5913 | -35.6479 | -38.2189
27.5974 | -35.6482 | -38.2246
27.6035 | -35.6485 | -38.2302
27.6096 | -35.6488 | -38.2356
27.6157 | -35.6491 | -38.241
27.6218 | -35.6494 | -38.2467
27.6279 | -35.6498 | -38.2523
27.634 | -35.6501 | -38.2577
27.6401 | -35.6504 | -38.2631
27.6462 | -35.6507 | -38.2683
27.6523 | -35.651 | -38.2735
27.6584 | -35.6513 | -38.279
27.6645 | -35.6516 | -38.2846
27.6706 | -35.6519 | -38.2898
27.6767 | -35.6523 | -38.2953
27.6828 | -35.6525 | -38.3006
27.6889 | -35.6529 | -38.3059
27.695 | -35.6532 | -38.3113
27.7011 | -35.6535 | -38.3168
27.7072 | -35.6539 | -38.3224
27.7133 | -35.6542 | -38.3279
27.7194 | -35.6545 | -38.3334
27.7256 | -35.6548 | -38.339
27.7317 | -35.6552 | -38.3444
27.7378 | -35.6555 | -38.35
27.7439 | -35.6558 | -38.3556
27.75 | -35.6562 | -38.3613
27.7561 | -35.6564 | -38.3665
27.7622 | -35.6567 | -38.3719
27.7683 | -35.657 | -38.3775
27.7744 | -35.6574 | -38.3831
27.7805 | -35.6577 | -38.3885
27.7866 | -35.6581 | -38.3941
27.7927 | -35.6584 | -38.3993
27.7988 | -35.6587 | -38.4049
27.8049 | -35.659 | -38.4103
27.811 | -35.6594 | -38.4155
27.8171 | -35.6597 | -38.4209
27.8232 | -35.66 | -38.4264
27.8293 | -35.6604 | -38.4317
27.8354 | -35.6607 | -38.4372
27.8415 | -35.6611 | -38.4426
27.8476 | -35.6614 | -38.4483
27.8537 | -35.6617 | -38.4539
27.8598 | -35.662 | -38.459
27.866 | -35.6623 | -38.4646
27.8721 | -35.6626 | -38.47
27.8782 | -35.663 | -38.4756
27.8843 | -35.6633 | -38.4812
27.8904 | -35.6636 | -38.4867
27.8965 | -35.6639 | -38.4923
27.9026 | -35.6642 | -38.4978
27.9087 | -35.6645 | -38.5034
27.9148 | -35.6648 | -38.5086
27.9209 | -35.6651 | -38.5142
27.927 | -35.6654 | -38.5193
27.9331 | -35.6656 | -38.5251
27.9392 | -35.666 | -38.5304
27.9453 | -35.6663 | -38.5355
27.9514 | -35.6666 | -38.5411
27.9575 | -35.6669 | -38.5466
27.9636 | -35.6672 | -38.552
27.9697 | -35.6675 | -38.5578
27.9758 | -35.6678 | -38.5634
27.9819 | -35.6681 | -38.5692
27.988 | -35.6684 | -38.5747
27.9941 | -35.6687 | -38.5801
28.0002 | -35.669 | -38.5854
28.0063 | -35.6692 | -38.5912
28.0125 | -35.6695 | -38.5966
28.0186 | -35.6698 | -38.6022
28.0247 | -35.6702 | -38.6076
28.0308 | -35.6706 | -38.6127
28.0369 | -35.6708 | -38.6178
28.043 | -35.6711 | -38.6234
28.0491 | -35.6714 | -38.6289
28.0552 | -35.6717 | -38.6347
28.0613 | -35.672 | -38.6404
28.0674 | -35.6723 | -38.6457
28.0735 | -35.6726 | -38.6506

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4876 | -43.6025
21.4908 | -35.5096 | -35.4974
16.4948 | -34.886 | -28.0062
12.6602 | -34.5238 | -21.6446
9.7171 | -34.3173 | -16.6949
7.4582 | -34.1641 | -12.6724
5.7244 | -34.0576 | -9.4229
4.3936 | -33.9883 | -6.7951
3.3722 | -33.9684 | -4.7324
2.5883 | -33.9504 | -3.1677
1.9866 | -33.9543 | -1.9157
1.5248 | -33.9434 | -0.91022
1.1703 | -33.9618 | -0.25958
1 | -33.9724 | 0.062401
0.89822 | -33.9717 | 0.33609
0.68941 | -33.9383 | 0.92478
0.52919 | -33.9388 | 1.3788
0.40612 | -33.9475 | 1.9835
0.31171 | -33.9438 | 2.4365
0.23928 | -33.9399 | 3.1675
0.18358 | -33.9395 | 3.594
0.14091 | -33.9507 | 4.4824
0.10819 | -33.9693 | 5.6468
0.083017 | -34.0092 | 7.1188
0.063705 | -34.0551 | 8.9934
0.048923 | -34.126 | 11.4797
0.037479 | -34.2327 | 14.6776
0.028801 | -34.4133 | 18.7056
0.022078 | -34.6968 | 24.1299
0.017023 | -35.165 | 30.5844
0.013018 | -35.9712 | 38.2917
0.010014 | -37.1006 | 46.2062

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003751 | -10.4891
0.007502 | -27.1957
0.011252 | -35.4718
0.015003 | -41.2047
0.018754 | -43.9334
0.022505 | -46.1335
0.026255 | -47.3703
0.030006 | -48.2282
0.033757 | -49.6251
0.037508 | -50.3372
0.041258 | -51.5401
0.045009 | -52.6509
0.04876 | -53.7038
0.052511 | -55.6864
0.056261 | -55.9091
0.060012 | -55.7914
0.063763 | -56.3312
0.067514 | -58.7312
0.071264 | -59.5376
0.075015 | -59.2906
0.078766 | -59.8965
0.082517 | -59.8922
0.086268 | -60.2133
0.090018 | -60.6043
0.093769 | -61.137
0.09752 | -61.1385
0.10127 | -61.5364
0.10502 | -62.4419
0.10877 | -62.4455
0.11252 | -62.4563
0.11627 | -63.7987
0.12002 | -63.2925
0.12377 | -64.0498
0.12753 | -64.8835
0.13128 | -64.3538
0.13503 | -64.7777
0.13878 | -64.7796
0.14253 | -65.1016
0.14628 | -66.0982
0.15003 | -66.7743
0.15378 | -66.6014
0.15753 | -66.1698
0.16128 | -66.0057
0.16503 | -67.8157
0.16878 | -68.0692
0.17253 | -68.0686
0.17629 | -68.5475
0.18004 | -68.9085
0.18379 | -69.3337
0.18754 | -70.4395
0.19129 | -70.6933
0.19504 | -71.8882
0.19879 | -69.7749
0.20254 | -69.2196
0.20629 | -69.8304
0.21004 | -70.3489
0.21379 | -70.4954
0.21754 | -72.4807
0.22129 | -72.0106
0.22505 | -71.3816
0.2288 | -72.4472
0.23255 | -73.292
0.2363 | -72.9521
0.24005 | -71.7721
0.2438 | -71.581
0.24755 | -72.5586
0.2513 | -72.1855
0.25505 | -72.5567
0.2588 | -72.5026
0.26255 | -73.1016
0.2663 | -72.6299
0.27005 | -72.7097
0.27381 | -73.1094
0.27756 | -74.4853
0.28131 | -73.9834
0.28506 | -73.3646
0.28881 | -73.2729
0.29256 | -73.6214
0.29631 | -73.8563
0.30006 | -75.1982
0.30381 | -75.8451
0.30756 | -75.7107
0.31131 | -75.7469
0.31506 | -76.2525
0.31882 | -75.3344
0.32257 | -75.9292
0.32632 | -76.1016
0.33007 | -76.5128
0.33382 | -77.4958
0.33757 | -76.8386
0.34132 | -75.373
0.34507 | -76.5768
0.34882 | -77.1193
0.35257 | -76.2612
0.35632 | -76.2993
0.36007 | -75.8596
0.36382 | -76.0759
0.36757 | -77.2019
0.37133 | -77.5809
0.37508 | -76.6204
0.37883 | -76.6329
0.38258 | -76.9696
0.38633 | -76.7016
0.39008 | -77.6771
0.39383 | -77.4807
0.39758 | -77.0575
0.40133 | -77.4564
0.40508 | -77.5625
0.40883 | -78.5765
0.41258 | -78.5345
0.41634 | -77.7763
0.42008 | -78.838
0.42384 | -79.1609
0.42759 | -78.4781
0.43134 | -77.8901
0.43509 | -78.537
0.43884 | -79.4787
0.44259 | -80.7097
0.44634 | -80.7457
0.45009 | -80.9262
0.45384 | -80.3504
0.45759 | -80.2665
0.46134 | -79.5574
0.46509 | -79.1982
0.46885 | -80.3284
0.4726 | -79.1184
0.47635 | -79.0525
0.4801 | -79.6784
0.48385 | -79.3204
0.4876 | -80.1781
0.49135 | -80.8354
0.4951 | -80.3444
0.49885 | -80.3262
0.5026 | -80.0745
0.50635 | -80.35
0.5101 | -80.3635
0.51385 | -80.6354
0.5176 | -79.9543
0.52136 | -79.9736
0.52511 | -80.2821
0.52886 | -81.2197
0.53261 | -81.1543
0.53636 | -81.6072
0.54011 | -80.7857
0.54386 | -80.1888
0.54761 | -80.3289
0.55136 | -80.6538
0.55511 | -80.4302
0.55886 | -80.1331
0.56261 | -80.2227
0.56637 | -80.7394
0.57012 | -81.3164
0.57387 | -81.392
0.57762 | -81.35
0.58137 | -81.6521
0.58512 | -81.8919
0.58887 | -81.7813
0.59262 | -81.3674
0.59637 | -81.12
0.60012 | -81.7655
0.60387 | -82.4808
0.60762 | -82.3972
0.61137 | -81.9844
0.61513 | -81.5195
0.61888 | -81.3379
0.62263 | -81.8116
0.62638 | -82.1223
0.63013 | -82.1867
0.63388 | -82.5356
0.63763 | -82.7487
0.64138 | -82.759
0.64513 | -82.9671
0.64888 | -83.2021
0.65263 | -83.3258
0.65638 | -83.4496
0.66013 | -83.2657
0.66388 | -83.2112
0.66764 | -83.3912
0.67139 | -83.2244
0.67514 | -83.0939
0.67889 | -83.3702
0.68264 | -83.4621
0.68639 | -83.5082
0.69014 | -83.733
0.69389 | -84.0148
0.69764 | -83.9627
0.70139 | -83.6623
0.70514 | -83.6155
0.70889 | -84.0472
0.71264 | -84.3624
0.7164 | -84.4267
0.72015 | -84.5532
0.7239 | -84.3739
0.72765 | -84.1025
0.7314 | -84.12
0.73515 | -84.2308
0.7389 | -84.4019
0.74265 | -84.2596
0.7464 | -83.7665
0.75015 | -83.5924
0.7539 | -83.4273
0.75765 | -83.1777
0.7614 | -83.3511
0.76516 | -83.5983
0.76891 | -83.8891
0.77266 | -84.3
0.77641 | -84.3516
0.78016 | -84.0924
0.78391 | -84.03
0.78766 | -84.0918
0.79141 | -84.2244
0.79516 | -84.7935
0.79891 | -85.1216
0.80266 | -84.9315
0.80641 | -84.8432
0.81017 | -84.813
0.81392 | -84.5762
0.81767 | -84.2326
0.82142 | -84.1859
0.82517 | -84.3317
0.82892 | -84.5845
0.83267 | -84.969
0.83642 | -85.3289
0.84017 | -85.3717
0.84392 | -85.7549
0.84767 | -86.1943
0.85142 | -85.9764
0.85517 | -85.3758
0.85892 | -85.2226
0.86267 | -85.5168
0.86643 | -85.5313
0.87018 | -85.4568
0.87393 | -85.6651
0.87768 | -85.7144
0.88143 | -85.6358
0.88518 | -85.8048
0.88893 | -85.854
0.89268 | -85.608
0.89643 | -85.6471
0.90018 | -86.0037
0.90393 | -86.4385
0.90768 | -86.9643
0.91143 | -86.8937
0.91519 | -86.5778
0.91894 | -86.1824
0.92269 | -85.6532
0.92644 | -84.9146
0.93019 | -84.7624
0.93394 | -85.331
0.93769 | -85.9928
0.94144 | -86.325
0.94519 | -86.4376
0.94894 | -86.4745
0.95269 | -86.7243
0.95644 | -86.9141
0.9602 | -86.9302
0.96395 | -86.8443
0.9677 | -86.8517
0.97145 | -86.5055
0.9752 | -85.872
0.97895 | -86.0301
0.9827 | -86.6268
0.98645 | -86.4428
0.9902 | -86.1089
0.99395 | -86.3961
0.9977 | -86.745
1.0015 | -86.408
1.0052 | -86.0204
1.009 | -86.4555
1.0127 | -87.4009
1.0165 | -87.4035
1.0202 | -86.7635
1.024 | -86.7178
1.0277 | -87.344
1.0315 | -88.348
1.0352 | -88.644
1.039 | -87.8131
1.0427 | -86.9864
1.0465 | -86.5311
1.0502 | -86.4692
1.054 | -86.5871
1.0577 | -86.704
1.0615 | -86.9765
1.0652 | -87.5702
1.069 | -88.28
1.0727 | -88.3541
1.0765 | -87.4434
1.0802 | -87.0416
1.084 | -87.0113
1.0877 | -87.183
1.0915 | -87.375
1.0952 | -87.652
1.099 | -87.9421
1.1027 | -88.1228
1.1065 | -88.1081
1.1102 | -87.8809
1.114 | -87.692
1.1177 | -87.709
1.1215 | -87.8078
1.1252 | -87.84
1.129 | -87.797
1.1327 | -87.8049
1.1365 | -87.9592
1.1402 | -88.1789
1.144 | -88.2577
1.1477 | -88.2398
1.1515 | -88.3356
1.1552 | -88.4778
1.159 | -88.4073
1.1627 | -88.1224
1.1665 | -87.9023
1.1702 | -87.925
1.174 | -88.1496
1.1777 | -88.4568
1.1815 | -88.7477
1.1852 | -88.9982
1.189 | -89.223
1.1927 | -89.3945
1.1965 | -89.4967
1.2002 | -89.4876
1.204 | -89.2242
1.2077 | -88.7918
1.2115 | -88.5873
1.2152 | -88.818
1.219 | -89.2669
1.2227 | -89.5083
1.2265 | -89.5087
1.2303 | -89.5142
1.234 | -89.4884
1.2378 | -89.3725
1.2415 | -89.197
1.2453 | -88.9084
1.249 | -88.646
1.2528 | -88.6636
1.2565 | -88.9165
1.2603 | -89.152
1.264 | -89.3376
1.2678 | -89.491
1.2715 | -89.3954
1.2753 | -89.0607
1.279 | -88.845
1.2828 | -88.9981
1.2865 | -89.5068
1.2903 | -90.2063
1.294 | -90.845
1.2978 | -91.0159
1.3015 | -90.5587
1.3053 | -90.0061
1.309 | -89.8963
1.3128 | -90.179
1.3165 | -90.4194
1.3203 | -90.3549
1.324 | -90.1464
1.3278 | -90.085
1.3315 | -90.2044
1.3353 | -90.3222
1.339 | -90.3086
1.3428 | -90.2182
1.3465 | -90.1652
1.3503 | -90.1246
1.354 | -89.9795
1.3578 | -89.833
1.3615 | -89.8071
1.3653 | -89.7715
1.369 | -89.6351
1.3728 | -89.4883
1.3765 | -89.3927
1.3803 | -89.4639
1.384 | -89.7451
1.3878 | -90.0491
1.3915 | -90.2017
1.3953 | -90.3151
1.399 | -90.4866
1.4028 | -90.5636
1.4065 | -90.4242
1.4103 | -90.2969
1.414 | -90.4329
1.4178 | -90.7032
1.4215 | -90.686
1.4253 | -90.32
1.429 | -90.0001
1.4328 | -89.9674
1.4365 | -90.0773
1.4403 | -90.15
1.444 | -90.2085
1.4478 | -90.1994
1.4515 | -90.0587
1.4553 | -89.9489
1.459 | -89.9998
1.4628 | -90.1634
1.4665 | -90.2736
1.4703 | -90.3144
1.474 | -90.4475
1.4778 | -90.6936
1.4816 | -90.9075
1.4853 | -90.9554
1.4891 | -90.8448
1.4928 | -90.7132
1.4966 | -90.6944
1.5003 | -90.889
1.5041 | -91.1809
1.5078 | -91.233
1.5116 | -90.953
1.5153 | -90.5643
1.5191 | -90.2291
1.5228 | -90.0124
1.5266 | -89.9848
1.5303 | -90.1739
1.5341 | -90.4257
1.5378 | -90.5629
1.5416 | -90.7256
1.5453 | -91.0234
1.5491 | -91.1565
1.5528 | -90.9705
1.5566 | -90.8744
1.5603 | -91.1512
1.5641 | -91.5672
1.5678 | -91.6755
1.5716 | -91.5041
1.5753 | -91.4462
1.5791 | -91.594
1.5828 | -91.7202
1.5866 | -91.6221
1.5903 | -91.3022
1.5941 | -90.9642
1.5978 | -90.7692
1.6016 | -90.7646
1.6053 | -90.951
1.6091 | -91.2064
1.6128 | -91.3175
1.6166 | -91.2911
1.6203 | -91.4083
1.6241 | -91.5524
1.6278 | -91.7231
1.6316 | -91.8314
1.6353 | -91.8121
1.6391 | -91.6758
1.6428 | -91.5089
1.6466 | -91.3931
1.6503 | -91.3657
1.6541 | -91.4547
1.6578 | -91.6505
1.6616 | -91.8463
1.6653 | -91.9199
1.6691 | -91.883
1.6728 | -91.8503
1.6766 | -91.894
1.6803 | -91.9984
1.6841 | -92.0825
1.6878 | -92.0759
1.6916 | -91.9871
1.6953 | -91.9192
1.6991 | -91.9198
1.7028 | -91.9093
1.7066 | -91.8626
1.7103 | -91.873
1.7141 | -92.0593
1.7178 | -92.416
1.7216 | -92.7703
1.7253 | -92.9042
1.7291 | -92.7364
1.7329 | -92.4035
1.7366 | -92.026
1.7404 | -91.5885
1.7441 | -91.1569
1.7479 | -90.8996
1.7516 | -90.935
1.7554 | -91.2369
1.7591 | -91.6707
1.7629 | -92.0717
1.7666 | -92.3278
1.7704 | -92.4225
1.7741 | -92.3832
1.7779 | -92.2608
1.7816 | -92.1585
1.7854 | -92.1948
1.7891 | -92.4075
1.7929 | -92.6913
1.7966 | -92.8717
1.8004 | -92.8792
1.8041 | -92.8249
1.8079 | -92.8157
1.8116 | -92.8369
1.8154 | -92.8088
1.8191 | -92.6632
1.8229 | -92.454
1.8266 | -92.3168
1.8304 | -92.2772
1.8341 | -92.2736
1.8379 | -92.2411
1.8416 | -92.1825
1.8454 | -92.1233
1.8491 | -92.1339
1.8529 | -92.3196
1.8566 | -92.6719
1.8604 | -93.0422
1.8641 | -93.2587
1.8679 | -93.2348
1.8716 | -93.04
1.8754 | -92.7694
1.8791 | -92.5261
1.8829 | -92.3582
1.8866 | -92.2794
1.8904 | -92.2811
1.8941 | -92.3518
1.8979 | -92.4734
1.9016 | -92.5759
1.9054 | -92.5948
1.9091 | -92.5645
1.9129 | -92.5735
1.9166 | -92.7146
1.9204 | -92.9617
1.9241 | -93.184
1.9279 | -93.2014
1.9316 | -92.9919
1.9354 | -92.6827
1.9391 | -92.4148
1.9429 | -92.2936
1.9466 | -92.3796
1.9504 | -92.6512
1.9541 | -93.0337
1.9579 | -93.3547
1.9616 | -93.5159
1.9654 | -93.6003
1.9691 | -93.7412
1.9729 | -93.9286
1.9767 | -94.0808
1.9804 | -94.1517
1.9842 | -94.1169
1.9879 | -93.8777
1.9917 | -93.4835
1.9954 | -93.1208
1.9992 | -92.932
2.0029 | -92.9605
2.0067 | -93.1856
2.0104 | -93.5374
2.0142 | -93.8031
2.0179 | -93.7448
2.0217 | -93.3326
2.0254 | -92.8292
2.0292 | -92.5304
2.0329 | -92.5656
2.0367 | -92.8809
2.0404 | -93.2776
2.0442 | -93.4977
2.0479 | -93.4425
2.0517 | -93.2384
2.0554 | -93.1062
2.0592 | -93.1138
2.0629 | -93.2128
2.0667 | -93.3221
2.0704 | -93.3906
2.0742 | -93.4522
2.0779 | -93.5309
2.0817 | -93.5851
2.0854 | -93.6101
2.0892 | -93.7147
2.0929 | -93.9969
2.0967 | -94.3818
2.1004 | -94.6299
2.1042 | -94.5595
2.1079 | -94.26
2.1117 | -93.9765
2.1154 | -93.8561
2.1192 | -93.8956
2.1229 | -94.0115
2.1267 | -94.0926
2.1304 | -94.0535
2.1342 | -93.8924
2.1379 | -93.7129
2.1417 | -93.6402
2.1454 | -93.7366
2.1492 | -93.9752
2.1529 | -94.2474
2.1567 | -94.4174
2.1604 | -94.4007
2.1642 | -94.1
2.1679 | -93.871
2.1717 | -93.6523
2.1754 | -93.5257
2.1792 | -93.5221
2.1829 | -93.6259
2.1867 | -93.7889
2.1904 | -93.9566
2.1942 | -94.1006
2.1979 | -94.2214
2.2017 | -94.3282
2.2054 | -94.408
2.2092 | -94.4813
2.213 | -94.5485
2.2167 | -94.5816
2.2205 | -94.5639
2.2242 | -94.451
2.228 | -94.2623
2.2317 | -94.0414
2.2355 | -93.8294
2.2392 | -93.6679
2.243 | -93.551
2.2467 | -93.4507
2.2505 | -93.3619
2.2542 | -93.3167
2.258 | -93.3574
2.2617 | -93.4937
2.2655 | -93.6883
2.2692 | -93.8428
2.273 | -93.9046
2.2767 | -93.8888
2.2805 | -93.8482
2.2842 | -93.8543
2.288 | -93.9041
2.2917 | -93.9544
2.2955 | -93.9818
2.2992 | -93.9911
2.303 | -94.0217
2.3067 | -94.107
2.3105 | -94.2526
2.3142 | -94.4375
2.318 | -94.6134
2.3217 | -94.7402
2.3255 | -94.7883
2.3292 | -94.7755
2.333 | -94.7275
2.3367 | -94.6806
2.3405 | -94.6446
2.3442 | -94.5833
2.348 | -94.4814
2.3517 | -94.359
2.3555 | -94.2666
2.3592 | -94.2503
2.363 | -94.2737
2.3667 | -94.2826
2.3705 | -94.2269
2.3742 | -94.131
2.378 | -94.0471
2.3817 | -94.0478
2.3855 | -94.1725
2.3892 | -94.3886
2.393 | -94.6421
2.3967 | -94.8644
2.4005 | -94.9986
2.4042 | -95.0358
2.408 | -94.9568
2.4117 | -94.7959
2.4155 | -94.6032
2.4192 | -94.4462
2.423 | -94.3395
2.4267 | -94.2952
2.4305 | -94.3202
2.4342 | -94.4053
2.438 | -94.5326
2.4417 | -94.665
2.4455 | -94.7419
2.4492 | -94.7242
2.453 | -94.6019
2.4567 | -94.4251
2.4605 | -94.2416
2.4643 | -94.0959
2.468 | -94.0058
2.4718 | -93.9776
2.4755 | -94.0276
2.4793 | -94.1576
2.483 | -94.3441
2.4868 | -94.5441
2.4905 | -94.7047
2.4943 | -94.7937
2.498 | -94.8029
2.5018 | -94.7531
2.5055 | -94.6795
2.5093 | -94.5927
2.513 | -94.5137
2.5168 | -94.4318
2.5205 | -94.3687
2.5243 | -94.3468
2.528 | -94.3962
2.5318 | -94.5306
2.5355 | -94.7327
2.5393 | -94.942
2.543 | -95.0694
2.5468 | -95.048
2.5505 | -94.9208
2.5543 | -94.7908
2.558 | -94.7396
2.5618 | -94.777
2.5655 | -94.856
2.5693 | -94.9065
2.573 | -94.8994
2.5768 | -94.8676
2.5805 | -94.8768
2.5843 | -94.9939
2.588 | -95.2166
2.5918 | -95.4584
2.5955 | -95.5953
2.5993 | -95.5645
2.603 | -95.4003
2.6068 | -95.1769
2.6105 | -94.9671
2.6143 | -94.8425
2.618 | -94.8293
2.6218 | -94.8917
2.6255 | -94.9652
2.6293 | -95.0095
2.633 | -95.021
2.6368 | -95.0202
2.6405 | -95.0316
2.6443 | -95.0409
2.648 | -95.0293
2.6518 | -94.9608
2.6555 | -94.8232
2.6593 | -94.6984
2.663 | -94.6712
2.6668 | -94.7624
2.6705 | -94.9402
2.6743 | -95.1671
2.678 | -95.4031
2.6818 | -95.6191
2.6855 | -95.8149
2.6893 | -95.9765
2.693 | -96.1085
2.6968 | -96.2065
2.7005 | -96.2482
2.7043 | -96.2021
2.7081 | -96.1568
2.7118 | -96.0766
2.7156 | -95.979
2.7193 | -95.8848
2.7231 | -95.8167
2.7268 | -95.7989
2.7306 | -95.8384
2.7343 | -95.8957
2.7381 | -95.9235
2.7418 | -95.8978
2.7456 | -95.8193
2.7493 | -95.7179
2.7531 | -95.648
2.7568 | -95.6228
2.7606 | -95.6585
2.7643 | -95.7327
2.7681 | -95.8133
2.7718 | -95.8802
2.7756 | -95.8866
2.7793 | -95.8119
2.7831 | -95.6798
2.7868 | -95.5175
2.7906 | -95.3829
2.7943 | -95.3248
2.7981 | -95.3568
2.8018 | -95.4384
2.8056 | -95.5256
2.8093 | -95.595
2.8131 | -95.6442
2.8168 | -95.6916
2.8206 | -95.7452
2.8243 | -95.793
2.8281 | -95.7861
2.8318 | -95.6898
2.8356 | -95.5186
2.8393 | -95.3425
2.8431 | -95.2339
2.8468 | -95.2396
2.8506 | -95.3688
2.8543 | -95.5837
2.8581 | -95.8377
2.8618 | -96.0779
2.8656 | -96.2601
2.8693 | -96.363
2.8731 | -96.3846
2.8768 | -96.3393
2.8806 | -96.2677
2.8843 | -96.2039
2.8881 | -96.1752
2.8918 | -96.2296
2.8956 | -96.3702
2.8993 | -96.5684
2.9031 | -96.7653
2.9068 | -96.8877
2.9106 | -96.8997
2.9143 | -96.7822
2.9181 | -96.5656
2.9218 | -96.3246
2.9256 | -96.1409
2.9293 | -96.0514
2.9331 | -96.0651
2.9368 | -96.1434
2.9406 | -96.2274
2.9443 | -96.2782
2.9481 | -96.2608
2.9518 | -96.1832
2.9556 | -96.0842
2.9594 | -95.9979
2.9631 | -95.9443
2.9669 | -95.9318
2.9706 | -95.9701
2.9744 | -96.0737
2.9781 | -96.2424
2.9819 | -96.4535
2.9856 | -96.6709
2.9894 | -96.849
2.9931 | -96.941
2.9969 | -96.9422
3.0006 | -96.8529
3.0044 | -96.6923
3.0081 | -96.528
3.0119 | -96.4
3.0156 | -96.3319
3.0194 | -96.301
3.0231 | -96.281
3.0269 | -96.2634
3.0306 | -96.2294
3.0344 | -96.1886
3.0381 | -96.153
3.0419 | -96.1485
3.0456 | -96.1909
3.0494 | -96.2846
3.0531 | -96.4351
3.0569 | -96.6005
3.0606 | -96.7464
3.0644 | -96.8308
3.0681 | -96.8392
3.0719 | -96.7631
3.0756 | -96.6208
3.0794 | -96.449
3.0831 | -96.2936
3.0869 | -96.1902
3.0906 | -96.1376
3.0944 | -96.1252
3.0981 | -96.1275
3.1019 | -96.108
3.1056 | -96.0755
3.1094 | -96.0546
3.1131 | -96.0702
3.1169 | -96.1462
3.1206 | -96.2873
3.1244 | -96.4608
3.1281 | -96.6053
3.1319 | -96.6719
3.1356 | -96.6576
3.1394 | -96.5903
3.1431 | -96.4925
3.1469 | -96.4018
3.1506 | -96.3536
3.1544 | -96.3647
3.1581 | -96.4406
3.1619 | -96.5731
3.1656 | -96.7292
3.1694 | -96.8627
3.1731 | -96.9475
3.1769 | -96.9695
3.1806 | -96.9251
3.1844 | -96.8121
3.1881 | -96.6484
3.1919 | -96.4614
3.1956 | -96.2844
3.1994 | -96.1428
3.2032 | -96.0636
3.2069 | -96.0636
3.2107 | -96.1391
3.2144 | -96.2728
3.2182 | -96.4306
3.2219 | -96.5925
3.2257 | -96.7522
3.2294 | -96.8972
3.2332 | -97.0336
3.2369 | -97.138
3.2407 | -97.2036
3.2444 | -97.1868
3.2482 | -97.1704
3.2519 | -97.1257
3.2557 | -97.056
3.2594 | -96.9618
3.2632 | -96.8602
3.2669 | -96.7818
3.2707 | -96.7456
3.2744 | -96.7566
3.2782 | -96.8185
3.2819 | -96.9261
3.2857 | -97.0556
3.2894 | -97.1853
3.2932 | -97.3015
3.2969 | -97.392
3.3007 | -97.4441
3.3044 | -97.4639
3.3082 | -97.4619
3.3119 | -97.4406
3.3157 | -97.4149
3.3194 | -97.3903
3.3232 | -97.3617
3.3269 | -97.3303
3.3307 | -97.3019
3.3344 | -97.2668
3.3382 | -97.2392
3.3419 | -97.2323
3.3457 | -97.2642
3.3494 | -97.32
3.3532 | -97.3889
3.3569 | -97.4491
3.3607 | -97.4587
3.3644 | -97.4116
3.3682 | -97.3377
3.3719 | -97.258
3.3757 | -97.2074
3.3794 | -97.2065
3.3832 | -97.2561
3.3869 | -97.3287
3.3907 | -97.4237
3.3944 | -97.5109
3.3982 | -97.5787
3.4019 | -97.6233
3.4057 | -97.6429
3.4094 | -97.6488
3.4132 | -97.6452
3.4169 | -97.6396
3.4207 | -97.6389
3.4244 | -97.6393
3.4282 | -97.6526
3.4319 | -97.6697
3.4357 | -97.6866
3.4394 | -97.6982
3.4432 | -97.7131
3.447 | -97.7175
3.4507 | -97.7192
3.4545 | -97.7064
3.4582 | -97.6572
3.462 | -97.5843
3.4657 | -97.4845
3.4695 | -97.37
3.4732 | -97.2675
3.477 | -97.1889
3.4807 | -97.1436
3.4845 | -97.1322
3.4882 | -97.1401
3.492 | -97.1575
3.4957 | -97.1759
3.4995 | -97.1673
3.5032 | -97.1322
3.507 | -97.0753
3.5107 | -97.0121
3.5145 | -96.9525
3.5182 | -96.8982
3.522 | -96.8828
3.5257 | -96.9218
3.5295 | -96.9981
3.5332 | -97.1073
3.537 | -97.2279
3.5407 | -97.3448
3.5445 | -97.4535
3.5482 | -97.5373
3.552 | -97.5827
3.5557 | -97.595
3.5595 | -97.5738
3.5632 | -97.5249
3.567 | -97.4775
3.5707 | -97.4391
3.5745 | -97.4113
3.5782 | -97.3856
3.582 | -97.3678
3.5857 | -97.3852
3.5895 | -97.445
3.5932 | -97.5638
3.597 | -97.7235
3.6007 | -97.9087
3.6045 | -98.0939
3.6082 | -98.2495
3.612 | -98.3457
3.6157 | -98.3644
3.6195 | -98.3146
3.6232 | -98.2192
3.627 | -98.1399
3.6307 | -98.0998
3.6345 | -98.1066
3.6382 | -98.1637
3.642 | -98.2337
3.6457 | -98.2969
3.6495 | -98.3261
3.6532 | -98.3241
3.657 | -98.2783
3.6607 | -98.2017
3.6645 | -98.1226
3.6682 | -98.0532
3.672 | -98.0126
3.6757 | -97.9975
3.6795 | -97.9973
3.6832 | -98.0059
3.687 | -98.021
3.6907 | -98.0222
3.6945 | -98.0045
3.6983 | -97.9827
3.702 | -97.9619
3.7058 | -97.9535
3.7095 | -97.9555
3.7133 | -97.9714
3.717 | -97.9864
3.7208 | -97.9977
3.7245 | -98.0154
3.7283 | -98.0279
3.732 | -98.0462
3.7358 | -98.0795
3.7395 | -98.1202
3.7433 | -98.1735
3.747 | -98.2411
3.7508 | -98.3174
3.7545 | -98.3851
3.7583 | -98.4343
3.762 | -98.4688
3.7658 | -98.482
3.7695 | -98.4826
3.7733 | -98.4792
3.777 | -98.4696
3.7808 | -98.4684
3.7845 | -98.5214
3.7883 | -98.5675
3.792 | -98.626
3.7958 | -98.6863
3.7995 | -98.7443
3.8033 | -98.7791
3.807 | -98.7839
3.8108 | -98.7494
3.8145 | -98.6864
3.8183 | -98.5816
3.822 | -98.4374
3.8258 | -98.2776
3.8295 | -98.1055
3.8333 | -97.9431
3.837 | -97.8088
3.8408 | -97.7061
3.8445 | -97.6444
3.8483 | -97.6338
3.852 | -97.6718
3.8558 | -97.7503
3.8595 | -97.8484
3.8633 | -97.948
3.867 | -98.0323
3.8708 | -98.0791
3.8745 | -98.0843
3.8783 | -98.0544
3.882 | -98.0129
3.8858 | -97.9801
3.8895 | -97.9658
3.8933 | -97.9791
3.897 | -98.0123
3.9008 | -98.0532
3.9045 | -98.0967
3.9083 | -98.1299
3.912 | -98.1349
3.9158 | -98.117
3.9195 | -98.0938
3.9233 | -98.0807
3.927 | -98.0764
3.9308 | -98.0933
3.9345 | -98.1413
3.9383 | -98.2114
3.9421 | -98.2991
3.9458 | -98.3847
3.9496 | -98.4506
3.9533 | -98.4908
3.9571 | -98.5122
3.9608 | -98.5153
3.9646 | -98.5048
3.9683 | -98.4731
3.9721 | -98.4199
3.9758 | -98.3615
3.9796 | -98.2993
3.9833 | -98.237
3.9871 | -98.1953
3.9908 | -98.1912
3.9946 | -98.2332
3.9983 | -98.3235
4.0021 | -98.4685
4.0058 | -98.6484
4.0096 | -98.8472
4.0133 | -99.0237
4.0171 | -99.1435
4.0208 | -99.1871
4.0246 | -99.1539
4.0283 | -99.0855
4.0321 | -99.0104
4.0358 | -98.9521
4.0396 | -98.9187
4.0433 | -98.9089
4.0471 | -98.9228
4.0508 | -98.9517
4.0546 | -98.9973
4.0583 | -99.0545
4.0621 | -99.11
4.0658 | -99.143
4.0696 | -99.1496
4.0733 | -99.1274
4.0771 | -99.0854
4.0808 | -99.0361
4.0846 | -98.991
4.0883 | -98.9672
4.0921 | -98.9468
4.0958 | -98.9499
4.0996 | -98.968
4.1033 | -98.9822
4.1071 | -98.996
4.1108 | -99.0101
4.1146 | -99.0204
4.1183 | -99.0283
4.1221 | -99.0341
4.1258 | -99.0489
4.1296 | -99.0758
4.1333 | -99.1057
4.1371 | -99.1321
4.1408 | -99.1589
4.1446 | -99.189
4.1483 | -99.225
4.1521 | -99.2474
4.1558 | -99.2546
4.1596 | -99.2379
4.1633 | -99.2034
4.1671 | -99.1656
4.1708 | -99.1163
4.1746 | -99.0734
4.1783 | -99.0318
4.1821 | -98.9972
4.1859 | -98.9728
4.1896 | -98.945
4.1934 | -98.9126
4.1971 | -98.8812
4.2009 | -98.8686
4.2046 | -98.8862
4.2084 | -98.9419
4.2121 | -99.0378
4.2159 | -99.1717
4.2196 | -99.3251
4.2234 | -99.4837
4.2271 | -99.6334
4.2309 | -99.7495
4.2346 | -99.8325
4.2384 | -99.8693
4.2421 | -99.8652
4.2459 | -99.8146
4.2496 | -99.7262
4.2534 | -99.6164
4.2571 | -99.5049
4.2609 | -99.4002
4.2646 | -99.3015
4.2684 | -99.2155
4.2721 | -99.1399
4.2759 | -99.0808
4.2796 | -99.047
4.2834 | -99.0353
4.2871 | -99.0472
4.2909 | -99.0967
4.2946 | -99.1856
4.2984 | -99.2945
4.3021 | -99.4064
4.3059 | -99.5112
4.3096 | -99.6039
4.3134 | -99.6784
4.3171 | -99.7206
4.3209 | -99.7042
4.3246 | -99.6391
4.3284 | -99.5027
4.3321 | -99.4038
4.3359 | -99.2966
4.3396 | -99.2047
4.3434 | -99.1282
4.3471 | -99.0772
4.3509 | -99.0677
4.3546 | -99.0911
4.3584 | -99.1352
4.3621 | -99.1992
4.3659 | -99.2758
4.3696 | -99.3558
4.3734 | -99.4296
4.3771 | -99.471
4.3809 | -99.4783
4.3846 | -99.4687
4.3884 | -99.449
4.3921 | -99.424
4.3959 | -99.4013
4.3996 | -99.3876
4.4034 | -99.3918
4.4071 | -99.4142
4.4109 | -99.4511
4.4146 | -99.5057
4.4184 | -99.5808
4.4221 | -99.6623
4.4259 | -99.7487
4.4296 | -99.84
4.4334 | -99.9219
4.4372 | -99.9863
4.4409 | -100.0295
4.4447 | -100.0493
4.4484 | -100.051
4.4522 | -100.0346
4.4559 | -99.9978
4.4597 | -99.9447
4.4634 | -99.8724
4.4672 | -99.7888
4.4709 | -99.7057
4.4747 | -99.6273
4.4784 | -99.5627
4.4822 | -99.5201
4.4859 | -99.5091
4.4897 | -99.5138
4.4934 | -99.5383
4.4972 | -99.5773
4.5009 | -99.6152
4.5047 | -99.6456
4.5084 | -99.6651
4.5122 | -99.6757
4.5159 | -99.686
4.5197 | -99.7022
4.5234 | -99.7207
4.5272 | -99.7368
4.5309 | -99.7536
4.5347 | -99.7692
4.5384 | -99.7814
4.5422 | -99.7937
4.5459 | -99.8218
4.5497 | -99.8579
4.5534 | -99.8999
4.5572 | -99.9532
4.5609 | -100.0195
4.5647 | -100.081
4.5684 | -100.1358
4.5722 | -100.1756
4.5759 | -100.2139
4.5797 | -100.2589
4.5834 | -100.306
4.5872 | -100.3508
4.5909 | -100.393
4.5947 | -100.4228
4.5984 | -100.4373
4.6022 | -100.4292
4.6059 | -100.4022
4.6097 | -100.3651
4.6134 | -100.317
4.6172 | -100.2666
4.6209 | -100.2233
4.6247 | -100.1953
4.6284 | -100.1714
4.6322 | -100.154
4.6359 | -100.1357
4.6397 | -100.111
4.6434 | -100.0774
4.6472 | -100.0316
4.6509 | -99.9844
4.6547 | -99.9531
4.6584 | -99.9393
4.6622 | -99.9459
4.6659 | -99.9684
4.6697 | -99.9992
4.6734 | -100.0375
4.6772 | -100.075
4.681 | -100.1121
4.6847 | -100.1567
4.6885 | -100.2109
4.6922 | -100.2875
4.696 | -100.3716
4.6997 | -100.4618
4.7035 | -100.5559
4.7072 | -100.6447
4.711 | -100.7091
4.7147 | -100.7434
4.7185 | -100.74
4.7222 | -100.6781
4.726 | -100.5738
4.7297 | -100.4446
4.7335 | -100.2948
4.7372 | -100.148
4.741 | -100.006
4.7447 | -99.8819
4.7485 | -99.7871
4.7522 | -99.7191
4.756 | -99.6604
4.7597 | -99.6073
4.7635 | -99.5583
4.7672 | -99.5161
4.771 | -99.4785
4.7747 | -99.4494
4.7785 | -99.439
4.7822 | -99.4577
4.786 | -99.5051
4.7897 | -99.5813
4.7935 | -99.7008
4.7972 | -99.8453
4.801 | -99.9985
4.8047 | -100.1524
4.8085 | -100.2792
4.8122 | -100.3704
4.816 | -100.4085
4.8197 | -100.3932
4.8235 | -100.3295
4.8272 | -100.2467
4.831 | -100.1552
4.8347 | -100.0686
4.8385 | -100.0017
4.8422 | -99.9709
4.846 | -99.9738
4.8497 | -100.0094
4.8535 | -100.0749
4.8572 | -100.1636
4.861 | -100.2629
4.8647 | -100.3557
4.8685 | -100.4168
4.8722 | -100.4861
4.876 | -100.548
4.8797 | -100.5961
4.8835 | -100.6281
4.8872 | -100.6339
4.891 | -100.6281
4.8947 | -100.6123
4.8985 | -100.586
4.9022 | -100.5519
4.906 | -100.5125
4.9097 | -100.4685
4.9135 | -100.4263
4.9172 | -100.3869
4.921 | -100.3591
4.9248 | -100.3433
4.9285 | -100.3436
4.9323 | -100.3646
4.936 | -100.3968
4.9398 | -100.4387
4.9435 | -100.4804
4.9473 | -100.5195
4.951 | -100.5581
4.9548 | -100.5973
4.9585 | -100.6379
4.9623 | -100.686
4.966 | -100.7352
4.9698 | -100.7723
4.9735 | -100.7981
4.9773 | -100.8187
4.981 | -100.8264
4.9848 | -100.818
4.9885 | -100.7977
4.9923 | -100.7646
4.996 | -100.7258
4.9998 | -100.682
5.0035 | -100.6386
5.0073 | -100.5974
5.011 | -100.5658
5.0148 | -100.5403
5.0185 | -100.5152
5.0223 | -100.497
5.026 | -100.4864
5.0298 | -100.4762
5.0335 | -100.4633
5.0373 | -100.451
5.041 | -100.4357
5.0448 | -100.4314
5.0485 | -100.4313
5.0523 | -100.4366
5.056 | -100.4444
5.0598 | -100.4568
5.0635 | -100.4694
5.0673 | -100.494
5.071 | -100.5222
5.0748 | -100.5484
5.0785 | -100.5722
5.0823 | -100.5959
5.086 | -100.615
5.0898 | -100.6262
5.0935 | -100.6343
5.0973 | -100.6328
5.101 | -100.6203
5.1048 | -100.5985
5.1085 | -100.577
5.1123 | -100.5666
5.116 | -100.5682
5.1198 | -100.5858
5.1235 | -100.6112
5.1273 | -100.6504
5.131 | -100.7012
5.1348 | -100.7586
5.1385 | -100.8163
5.1423 | -100.8843
5.146 | -100.9422
5.1498 | -100.9927
5.1535 | -101.0365
5.1573 | -101.0808
5.161 | -101.1089
5.1648 | -101.122
5.1685 | -101.1172
5.1723 | -101.091
5.1761 | -101.0523
5.1798 | -101.0045
5.1836 | -100.9518
5.1873 | -100.9076
5.1911 | -100.8744
5.1948 | -100.855
5.1986 | -100.8603
5.2023 | -100.8907
5.2061 | -100.9337
5.2098 | -100.9809
5.2136 | -101.0289
5.2173 | -101.0714
5.2211 | -101.098
5.2248 | -101.112
5.2286 | -101.1129
5.2323 | -101.096
5.2361 | -101.0598
5.2398 | -101.0105
5.2436 | -100.9542
5.2473 | -100.8908
5.2511 | -100.8287
5.2548 | -100.7614
5.2586 | -100.6926
5.2623 | -100.6363
5.2661 | -100.591
5.2698 | -100.5593
5.2736 | -100.544
5.2773 | -100.5477
5.2811 | -100.5695
5.2848 | -100.613
5.2886 | -100.6734
5.2923 | -100.7446
5.2961 | -100.8224
5.2998 | -100.9016
5.3036 | -100.9787
5.3073 | -101.042
5.3111 | -101.0896
5.3148 | -101.1189
5.3186 | -101.126
5.3223 | -101.1081
5.3261 | -101.0792
5.3298 | -101.044
5.3336 | -101.0066
5.3373 | -100.9647
5.3411 | -100.9226
5.3448 | -100.88
5.3486 | -100.8432
5.3523 | -100.8106
5.3561 | -100.7816
5.3598 | -100.7677
5.3636 | -100.76
5.3673 | -100.7592
5.3711 | -100.7604
5.3748 | -100.7614
5.3786 | -100.7627
5.3823 | -100.7521
5.3861 | -100.7372
5.3898 | -100.7218
5.3936 | -100.7079
5.3973 | -100.7028
5.4011 | -100.7099
5.4048 | -100.7266
5.4086 | -100.7886
5.4123 | -100.8318
5.4161 | -100.8851
5.4199 | -100.9465
5.4236 | -101.0096
5.4274 | -101.0722
5.4311 | -101.1246
5.4349 | -101.1584
5.4386 | -101.171
5.4424 | -101.1673
5.4461 | -101.1479
5.4499 | -101.1032
5.4536 | -101.0519
5.4574 | -101.0004
5.4611 | -100.9468
5.4649 | -100.9064
5.4686 | -100.8756
5.4724 | -100.8563
5.4761 | -100.8515
5.4799 | -100.8709
5.4836 | -100.9034
5.4874 | -100.952
5.4911 | -101.0094
5.4949 | -101.0827
5.4986 | -101.1589
5.5024 | -101.2391
5.5061 | -101.315
5.5099 | -101.3748
5.5136 | -101.4095
5.5174 | -101.4219
5.5211 | -101.4107
5.5249 | -101.3691
5.5286 | -101.2985
5.5324 | -101.223
5.5361 | -101.1547
5.5399 | -101.0854
5.5436 | -101.0198
5.5474 | -100.9745
5.5511 | -100.9469
5.5549 | -100.9336
5.5586 | -100.9256
5.5624 | -100.9264
5.5661 | -100.9326
5.5699 | -100.9494
5.5736 | -100.9728
5.5774 | -101
5.5811 | -101.021
5.5849 | -101.0417
5.5886 | -101.0571
5.5924 | -101.0693
5.5961 | -101.072
5.5999 | -101.0754
5.6036 | -101.0792
5.6074 | -101.0854
5.6111 | -101.1058
5.6149 | -101.1349
5.6186 | -101.1651
5.6224 | -101.1881
5.6261 | -101.2106
5.6299 | -101.2383
5.6336 | -101.2609
5.6374 | -101.271
5.6411 | -101.2847
5.6449 | -101.3085
5.6486 | -101.3382
5.6524 | -101.3702
5.6561 | -101.4019
5.6599 | -101.4366
5.6636 | -101.4766
5.6674 | -101.5143
5.6712 | -101.5494
5.6749 | -101.5782
5.6787 | -101.5966
5.6824 | -101.6156
5.6862 | -101.6298
5.6899 | -101.6303
5.6937 | -101.6177
5.6974 | -101.5905
5.7012 | -101.5523
5.7049 | -101.51
5.7087 | -101.4757
5.7124 | -101.4449
5.7162 | -101.4103
5.7199 | -101.385
5.7237 | -101.376
5.7274 | -101.3704
5.7312 | -101.3653
5.7349 | -101.3613
5.7387 | -101.3613
5.7424 | -101.3665
5.7462 | -101.3656
5.7499 | -101.3586
5.7537 | -101.3509
5.7574 | -101.3457
5.7612 | -101.3412
5.7649 | -101.3441
5.7687 | -101.3492
5.7724 | -101.3558
5.7762 | -101.3746
5.7799 | -101.4033
5.7837 | -101.4439
5.7874 | -101.4942
5.7912 | -101.5496
5.7949 | -101.6015
5.7987 | -101.6491
5.8024 | -101.7028
5.8062 | -101.7564
5.8099 | -101.8071
5.8137 | -101.8528
5.8174 | -101.8928
5.8212 | -101.9363
5.8249 | -101.9697
5.8287 | -101.9924
5.8324 | -102.0157
5.8362 | -102.0331
5.8399 | -102.0352
5.8437 | -102.0236
5.8474 | -102.0012
5.8512 | -101.9633
5.8549 | -101.915
5.8587 | -101.8636
5.8624 | -101.7989
5.8662 | -101.7241
5.8699 | -101.6442
5.8737 | -101.5692
5.8774 | -101.5043
5.8812 | -101.4527
5.8849 | -101.4091
5.8887 | -101.3645
5.8924 | -101.325
5.8962 | -101.3052
5.8999 | -101.2913
5.9037 | -101.2784
5.9074 | -101.2658
5.9112 | -101.256
5.915 | -101.2431
5.9187 | -101.2224
5.9225 | -101.1966
5.9262 | -101.1645
5.93 | -101.1342
5.9337 | -101.1133
5.9375 | -101.1038
5.9412 | -101.0948
5.945 | -101.0891
5.9487 | -101.0949
5.9525 | -101.0924
5.9562 | -101.0878
5.96 | -101.0721
5.9637 | -101.045
5.9675 | -101.013
5.9712 | -100.9797
5.975 | -100.9533
5.9787 | -100.9386
5.9825 | -100.9345
5.9862 | -100.9377
5.99 | -100.9599
5.9937 | -100.9971
5.9975 | -101.0404
6.0012 | -101.0944
6.005 | -101.1544
6.0087 | -101.2184
6.0125 | -101.2913
6.0162 | -101.3733
6.02 | -101.4547
6.0237 | -101.5362
6.0275 | -101.6137
6.0312 | -101.6935
6.035 | -101.7637
6.0387 | -101.8169
6.0425 | -101.8581
6.0462 | -101.8896
6.05 | -101.9083
6.0537 | -101.9163
6.0575 | -101.9077
6.0612 | -101.8827
6.065 | -101.8634
6.0687 | -101.8401
6.0725 | -101.8104
6.0762 | -101.7748
6.08 | -101.7436
6.0837 | -101.7193
6.0875 | -101.7034
6.0912 | -101.6945
6.095 | -101.6952
6.0987 | -101.7068
6.1025 | -101.7284
6.1062 | -101.7556
6.11 | -101.7886
6.1137 | -101.8303
6.1175 | -101.8753
6.1212 | -101.922
6.125 | -101.9662
6.1287 | -101.9928
6.1325 | -102.0092
6.1362 | -102.0142
6.14 | -102.011
6.1437 | -101.9919
6.1475 | -101.9572
6.1512 | -101.9202
6.155 | -101.8829
6.1588 | -101.8329
6.1625 | -101.7891
6.1663 | -101.7614
6.17 | -101.748
6.1738 | -101.7455
6.1775 | -101.751
6.1813 | -101.7722
6.185 | -101.8105
6.1888 | -101.8492
6.1925 | -101.8862
6.1963 | -101.9202
6.2 | -101.9476
6.2038 | -101.9654
6.2075 | -101.969
6.2113 | -101.9616
6.215 | -101.9497
6.2188 | -101.9338
6.2225 | -101.9278
6.2263 | -101.9183
6.23 | -101.9197
6.2338 | -101.9246
6.2375 | -101.9345
6.2413 | -101.9541
6.245 | -101.977
6.2488 | -101.9996
6.2525 | -102.0293
6.2563 | -102.0617
6.26 | -102.0991
6.2638 | -102.1428
6.2675 | -102.1842
6.2713 | -102.2346
6.275 | -102.2803
6.2788 | -102.3176
6.2825 | -102.3548
6.2863 | -102.3838
6.29 | -102.393
6.2938 | -102.3847
6.2975 | -102.3627
6.3013 | -102.3257
6.305 | -102.2843
6.3088 | -102.2296
6.3125 | -102.1695
6.3163 | -102.108
6.32 | -102.0566
6.3238 | -102.0114
6.3275 | -101.9754
6.3313 | -101.956
6.335 | -101.9456
6.3388 | -101.942
6.3425 | -101.9495
6.3463 | -101.964
6.35 | -101.9798
6.3538 | -101.9921
6.3575 | -102.0073
6.3613 | -102.0245
6.365 | -102.0357
6.3688 | -102.0454
6.3725 | -102.0526
6.3763 | -102.0576
6.38 | -102.0576
6.3838 | -102.0608
6.3875 | -102.0684
6.3913 | -102.0793
6.395 | -102.0919
6.3988 | -102.1076
6.4026 | -102.1315
6.4063 | -102.1497
6.4101 | -102.1667
6.4138 | -102.1802
6.4176 | -102.1919
6.4213 | -102.2018
6.4251 | -102.2087
6.4288 | -102.2176
6.4326 | -102.2291
6.4363 | -102.2451
6.4401 | -102.267
6.4438 | -102.2935
6.4476 | -102.3202
6.4513 | -102.3457
6.4551 | -102.3667
6.4588 | -102.3819
6.4626 | -102.3838
6.4663 | -102.3815
6.4701 | -102.3721
6.4738 | -102.3549
6.4776 | -102.3248
6.4813 | -102.2866
6.4851 | -102.2525
6.4888 | -102.2221
6.4926 | -102.2292
6.4963 | -102.2211
6.5001 | -102.2252
6.5038 | -102.2491
6.5076 | -102.2844
6.5113 | -102.334
6.5151 | -102.3941
6.5188 | -102.461
6.5226 | -102.5275
6.5263 | -102.5919
6.5301 | -102.649
6.5338 | -102.689
6.5376 | -102.7137
6.5413 | -102.7263
6.5451 | -102.7227
6.5488 | -102.6998
6.5526 | -102.6686
6.5563 | -102.6329
6.5601 | -102.5947
6.5638 | -102.5492
6.5676 | -102.5039
6.5713 | -102.4665
6.5751 | -102.4427
6.5788 | -102.4301
6.5826 | -102.4261
6.5863 | -102.437
6.5901 | -102.4543
6.5938 | -102.4814
6.5976 | -102.5147
6.6013 | -102.5488
6.6051 | -102.5813
6.6088 | -102.6096
6.6126 | -102.6233
6.6163 | -102.6281
6.6201 | -102.6244
6.6238 | -102.6041
6.6276 | -102.5712
6.6313 | -102.5266
6.6351 | -102.4782
6.6388 | -102.4267
6.6426 | -102.3692
6.6463 | -102.3072
6.6501 | -102.252
6.6539 | -102.1978
6.6576 | -102.1431
6.6614 | -102.0997
6.6651 | -102.0685
6.6689 | -102.0431
6.6726 | -102.0277
6.6764 | -102.0169
6.6801 | -102.0086
6.6839 | -102.0078
6.6876 | -102.0097
6.6914 | -102.0142
6.6951 | -102.0183
6.6989 | -102.0223
6.7026 | -102.026
6.7064 | -102.0196
6.7101 | -102.0192
6.7139 | -102.0293
6.7176 | -102.0373
6.7214 | -102.0439
6.7251 | -102.0483
6.7289 | -102.0493
6.7326 | -102.0466
6.7364 | -102.038
6.7401 | -102.0277
6.7439 | -102.0206
6.7476 | -102.0224
6.7514 | -102.0256
6.7551 | -102.0387
6.7589 | -102.0664
6.7626 | -102.1029
6.7664 | -102.1493
6.7701 | -102.2101
6.7739 | -102.2752
6.7776 | -102.3428
6.7814 | -102.4124
6.7851 | -102.4735
6.7889 | -102.5215
6.7926 | -102.5591
6.7964 | -102.5882
6.8001 | -102.6014
6.8039 | -102.6041
6.8076 | -102.5996
6.8114 | -102.5751
6.8151 | -102.5439
6.8189 | -102.5073
6.8226 | -102.4644
6.8264 | -102.424
6.8301 | -102.3802
6.8339 | -102.3397
6.8376 | -102.3025
6.8414 | -102.2752
6.8451 | -102.2614
6.8489 | -102.2511
6.8526 | -102.254
6.8564 | -102.2655
6.8601 | -102.2808
6.8639 | -102.3006
6.8676 | -102.3203
6.8714 | -102.3336
6.8751 | -102.341
6.8789 | -102.345
6.8826 | -102.3414
6.8864 | -102.3324
6.8902 | -102.3283
6.8939 | -102.3282
6.8977 | -102.3325
6.9014 | -102.3436
6.9052 | -102.3604
6.9089 | -102.3915
6.9127 | -102.434
6.9164 | -102.4747
6.9202 | -102.5212
6.9239 | -102.5675
6.9277 | -102.6079
6.9314 | -102.6423
6.9352 | -102.6674
6.9389 | -102.6838
6.9427 | -102.6941
6.9464 | -102.6991
6.9502 | -102.7043
6.9539 | -102.7046
6.9577 | -102.7087
6.9614 | -102.7122
6.9652 | -102.7138
6.9689 | -102.7163
6.9727 | -102.7193
6.9764 | -102.7199
6.9802 | -102.7157
6.9839 | -102.7134
6.9877 | -102.7128
6.9914 | -102.7054
6.9952 | -102.6984
6.9989 | -102.6975
7.0027 | -102.7105
7.0064 | -102.7291
7.0102 | -102.7473
7.0139 | -102.7686
7.0177 | -102.7963
7.0214 | -102.8295
7.0252 | -102.8646
7.0289 | -102.8982
7.0327 | -102.9233
7.0364 | -102.9457
7.0402 | -102.9681
7.0439 | -102.9897
7.0477 | -103.0043
7.0514 | -103.0238
7.0552 | -103.0425
7.0589 | -103.0507
7.0627 | -103.0534
7.0664 | -103.0568
7.0702 | -103.0607
7.0739 | -103.0594
7.0777 | -103.0552
7.0814 | -103.0463
7.0852 | -103.0324
7.0889 | -103.0115
7.0927 | -102.9831
7.0964 | -102.9561
7.1002 | -102.9218
7.1039 | -102.8872
7.1077 | -102.8517
7.1114 | -102.8229
7.1152 | -102.8009
7.1189 | -102.781
7.1227 | -102.7625
7.1264 | -102.7367
7.1302 | -102.7204
7.1339 | -102.709
7.1377 | -102.7024
7.1415 | -102.697
7.1452 | -102.6981
7.149 | -102.6963
7.1527 | -102.6997
7.1565 | -102.7068
7.1602 | -102.7161
7.164 | -102.7414
7.1677 | -102.7785
7.1715 | -102.8199
7.1752 | -102.8642
7.179 | -102.9139
7.1827 | -102.9709
7.1865 | -103.0332
7.1902 | -103.0995
7.194 | -103.16
7.1977 | -103.2127
7.2015 | -103.2592
7.2052 | -103.2989
7.209 | -103.3339
7.2127 | -103.3556
7.2165 | -103.3641
7.2202 | -103.3625
7.224 | -103.3573
7.2277 | -103.3519
7.2315 | -103.3467
7.2352 | -103.3417
7.239 | -103.3315
7.2427 | -103.3225
7.2465 | -103.317
7.2502 | -103.32
7.254 | -103.3246
7.2577 | -103.3372
7.2615 | -103.347
7.2652 | -103.3598
7.269 | -103.3701
7.2727 | -103.3744
7.2765 | -103.3773
7.2802 | -103.3756
7.284 | -103.3745
7.2877 | -103.3671
7.2915 | -103.3523
7.2952 | -103.3327
7.299 | -103.3059
7.3027 | -103.2803
7.3065 | -103.2509
7.3102 | -103.2171
7.314 | -103.1857
7.3177 | -103.1569
7.3215 | -103.1283
7.3252 | -103.1018
7.329 | -103.0727
7.3327 | -103.0428
7.3365 | -103.0123
7.3402 | -102.9847
7.344 | -102.9639
7.3477 | -102.939
7.3515 | -102.9208
7.3552 | -102.9129
7.359 | -102.9132
7.3627 | -102.9205
7.3665 | -102.9336
7.3702 | -102.9526
7.374 | -102.9731
7.3777 | -102.9998
7.3815 | -103.0294
7.3853 | -103.0568
7.389 | -103.0771
7.3928 | -103.0904
7.3965 | -103.0997
7.4003 | -103.1068
7.404 | -103.106
7.4078 | -103.0986
7.4115 | -103.0873
7.4153 | -103.0663
7.419 | -103.0364
7.4228 | -102.9982
7.4265 | -102.9577
7.4303 | -102.9225
7.434 | -102.8917
7.4378 | -102.8658
7.4415 | -102.8452
7.4453 | -102.835
7.449 | -102.8325
7.4528 | -102.8376
7.4565 | -102.8555
7.4603 | -102.882
7.464 | -102.9138
7.4678 | -102.9513
7.4715 | -102.9947
7.4753 | -103.0389
7.479 | -103.0806
7.4828 | -103.1226
7.4865 | -103.1572
7.4903 | -103.1826
7.494 | -103.1982
7.4978 | -103.208
7.5015 | -103.208
7.5053 | -103.1958
7.509 | -103.1755
7.5128 | -103.1512
7.5165 | -103.1272
7.5203 | -103.0947
7.524 | -103.064
7.5278 | -103.0305
7.5315 | -102.9974
7.5353 | -102.9684
7.539 | -102.9432
7.5428 | -102.9186
7.5465 | -102.9016
7.5503 | -102.8904
7.554 | -102.8855
7.5578 | -102.8864
7.5615 | -102.8982
7.5653 | -102.9201
7.569 | -102.9517
7.5728 | -103.0098
7.5765 | -103.0543
7.5803 | -103.1032
7.584 | -103.1463
7.5878 | -103.1917
7.5915 | -103.2393
7.5953 | -103.2804
7.599 | -103.3096
7.6028 | -103.3275
7.6065 | -103.3456
7.6103 | -103.3537
7.614 | -103.3511
7.6178 | -103.3398
7.6215 | -103.3186
7.6253 | -103.2881
7.6291 | -103.2571
7.6328 | -103.2308
7.6366 | -103.2062
7.6403 | -103.1851
7.6441 | -103.1633
7.6478 | -103.1499
7.6516 | -103.1476
7.6553 | -103.1572
7.6591 | -103.1734
7.6628 | -103.1937
7.6666 | -103.2156
7.6703 | -103.248
7.6741 | -103.2843
7.6778 | -103.3145
7.6816 | -103.3405
7.6853 | -103.3598
7.6891 | -103.3758
7.6928 | -103.3833
7.6966 | -103.3853
7.7003 | -103.3826
7.7041 | -103.3689
7.7078 | -103.3451
7.7116 | -103.3185
7.7153 | -103.2964
7.7191 | -103.2801
7.7228 | -103.2594
7.7266 | -103.2498
7.7303 | -103.2498
7.7341 | -103.2614
7.7378 | -103.2837
7.7416 | -103.3184
7.7453 | -103.3557
7.7491 | -103.3911
7.7528 | -103.4293
7.7566 | -103.472
7.7603 | -103.516
7.7641 | -103.5503
7.7678 | -103.5779
7.7716 | -103.6028
7.7753 | -103.6196
7.7791 | -103.6367
7.7828 | -103.6497
7.7866 | -103.6507
7.7903 | -103.6457
7.7941 | -103.6442
7.7978 | -103.6451
7.8016 | -103.6481
7.8053 | -103.6508
7.8091 | -103.6524
7.8128 | -103.6567
7.8166 | -103.6668
7.8203 | -103.6747
7.8241 | -103.6804
7.8278 | -103.6884
7.8316 | -103.6961
7.8353 | -103.7018
7.8391 | -103.7027
7.8428 | -103.6902
7.8466 | -103.6729
7.8503 | -103.6529
7.8541 | -103.6299
7.8578 | -103.5994
7.8616 | -103.5702
7.8653 | -103.5449
7.8691 | -103.5174
7.8728 | -103.4872
7.8766 | -103.4595
7.8804 | -103.4354
7.8841 | -103.4113
7.8879 | -103.385
7.8916 | -103.3599
7.8954 | -103.3343
7.8991 | -103.3074
7.9029 | -103.2813
7.9066 | -103.2538
7.9104 | -103.2251
7.9141 | -103.1937
7.9179 | -103.1676
7.9216 | -103.1429
7.9254 | -103.1244
7.9291 | -103.1143
7.9329 | -103.1105
7.9366 | -103.1153
7.9404 | -103.1276
7.9441 | -103.1401
7.9479 | -103.1617
7.9516 | -103.1935
7.9554 | -103.2366
7.9591 | -103.2829
7.9629 | -103.3293
7.9666 | -103.3699
7.9704 | -103.4065
7.9741 | -103.4443
7.9779 | -103.477
7.9816 | -103.506
7.9854 | -103.5303
7.9891 | -103.5511
7.9929 | -103.5663
7.9966 | -103.5808
8.0004 | -103.5929
8.0041 | -103.6052
8.0079 | -103.6198
8.0116 | -103.6372
8.0154 | -103.6535
8.0191 | -103.6748
8.0229 | -103.6961
8.0266 | -103.7158
8.0304 | -103.736
8.0341 | -103.7493
8.0379 | -103.7627
8.0416 | -103.773
8.0454 | -103.7784
8.0491 | -103.7818
8.0529 | -103.7832
8.0566 | -103.786
8.0604 | -103.7873
8.0641 | -103.7838
8.0679 | -103.7783
8.0716 | -103.7711
8.0754 | -103.7698
8.0791 | -103.7753
8.0829 | -103.7818
8.0866 | -103.7917
8.0904 | -103.8023
8.0941 | -103.8097
8.0979 | -103.812
8.1016 | -103.8177
8.1054 | -103.8241
8.1091 | -103.8267
8.1129 | -103.8236
8.1166 | -103.8082
8.1204 | -103.8028
8.1242 | -103.7936
8.1279 | -103.7806
8.1317 | -103.7582
8.1354 | -103.7379
8.1392 | -103.7137
8.1429 | -103.6847
8.1467 | -103.6626
8.1504 | -103.6442
8.1542 | -103.6275
8.1579 | -103.6107
8.1617 | -103.5995
8.1654 | -103.5933
8.1692 | -103.5936
8.1729 | -103.5971
8.1767 | -103.6019
8.1804 | -103.6089
8.1842 | -103.6116
8.1879 | -103.6143
8.1917 | -103.6208
8.1954 | -103.62
8.1992 | -103.6113
8.2029 | -103.6033
8.2067 | -103.5944
8.2104 | -103.5861
8.2142 | -103.5832
8.2179 | -103.581
8.2217 | -103.5791
8.2254 | -103.5806
8.2292 | -103.5905
8.2329 | -103.6042
8.2367 | -103.6133
8.2404 | -103.6229
8.2442 | -103.6362
8.2479 | -103.6504
8.2517 | -103.6662
8.2554 | -103.6827
8.2592 | -103.6945
8.2629 | -103.7012
8.2667 | -103.7123
8.2704 | -103.7187
8.2742 | -103.7199
8.2779 | -103.724
8.2817 | -103.7256
8.2854 | -103.7203
8.2892 | -103.71
8.2929 | -103.6959
8.2967 | -103.6827
8.3004 | -103.6695
8.3042 | -103.6551
8.3079 | -103.6399
8.3117 | -103.6216
8.3154 | -103.5983
8.3192 | -103.5802
8.3229 | -103.5632
8.3267 | -103.5469
8.3304 | -103.5289
8.3342 | -103.5148
8.3379 | -103.5066
8.3417 | -103.5075
8.3454 | -103.5078
8.3492 | -103.5076
8.3529 | -103.5145
8.3567 | -103.5297
8.3604 | -103.5548
8.3642 | -103.5861
8.368 | -103.6256
8.3717 | -103.6699
8.3755 | -103.7142
8.3792 | -103.7637
8.383 | -103.8208
8.3867 | -103.8796
8.3905 | -103.9311
8.3942 | -103.9775
8.398 | -104.0218
8.4017 | -104.0537
8.4055 | -104.0775
8.4092 | -104.0936
8.413 | -104.0991
8.4167 | -104.0922
8.4205 | -104.0727
8.4242 | -104.0466
8.428 | -104.017
8.4317 | -103.9851
8.4355 | -103.9539
8.4392 | -103.9251
8.443 | -103.8983
8.4467 | -103.879
8.4505 | -103.8622
8.4542 | -103.8503
8.458 | -103.8408
8.4617 | -103.8368
8.4655 | -103.8397
8.4692 | -103.8485
8.473 | -103.8664
8.4767 | -103.888
8.4805 | -103.9146
8.4842 | -103.944
8.488 | -103.9777
8.4917 | -104.0145
8.4955 | -104.0555
8.4992 | -104.1002
8.503 | -104.1444
8.5067 | -104.1783
8.5105 | -104.2093
8.5142 | -104.2409
8.518 | -104.2684
8.5217 | -104.2874
8.5255 | -104.3033
8.5292 | -104.3098
8.533 | -104.3015
8.5367 | -104.2876
8.5405 | -104.269
8.5442 | -104.2444
8.548 | -104.2115
8.5517 | -104.1735
8.5555 | -104.1335
8.5592 | -104.0924
8.563 | -104.0558
8.5667 | -104.0229
8.5705 | -103.9872
8.5742 | -103.962
8.578 | -103.9471
8.5817 | -103.9417
8.5855 | -103.9424
8.5892 | -103.9445
8.593 | -103.9522
8.5967 | -103.9673
8.6005 | -103.9796
8.6042 | -103.9924
8.608 | -104.0089
8.6117 | -104.0217
8.6155 | -104.0271
8.6193 | -104.0305
8.623 | -104.0291
8.6268 | -104.0172
8.6305 | -104.006
8.6343 | -103.9848
8.638 | -103.9621
8.6418 | -103.9353
8.6455 | -103.9076
8.6493 | -103.8759
8.653 | -103.842
8.6568 | -103.819
8.6605 | -103.8044
8.6643 | -103.7918
8.668 | -103.7777
8.6718 | -103.7641
8.6755 | -103.7538
8.6793 | -103.7471
8.683 | -103.7447
8.6868 | -103.7448
8.6905 | -103.7422
8.6943 | -103.7433
8.698 | -103.7532
8.7018 | -103.7615
8.7055 | -103.7694
8.7093 | -103.7764
8.713 | -103.7832
8.7168 | -103.7988
8.7205 | -103.8176
8.7243 | -103.8337
8.728 | -103.8542
8.7318 | -103.8753
8.7355 | -103.8921
8.7393 | -103.9081
8.743 | -103.9256
8.7468 | -103.9389
8.7505 | -103.9489
8.7543 | -103.9651
8.758 | -103.9767
8.7618 | -103.9864
8.7655 | -103.9943
8.7693 | -104.0081
8.773 | -104.023
8.7768 | -104.0363
8.7805 | -104.051
8.7843 | -104.0631
8.788 | -104.0655
8.7918 | -104.0659
8.7955 | -104.0646
8.7993 | -104.0631
8.803 | -104.0651
8.8068 | -104.068
8.8105 | -104.0688
8.8143 | -104.0611
8.818 | -104.0525
8.8218 | -104.0494
8.8255 | -104.042
8.8293 | -104.0333
8.833 | -104.0285
8.8368 | -104.0239
8.8405 | -104.0214
8.8443 | -104.0209
8.848 | -104.0165
8.8518 | -104.0176
8.8555 | -104.0182
8.8593 | -104.0192
8.863 | -104.0232
8.8668 | -104.0294
8.8706 | -104.0359
8.8743 | -104.0421
8.8781 | -104.0524
8.8818 | -104.0634
8.8856 | -104.0698
8.8893 | -104.0791
8.8931 | -104.0928
8.8968 | -104.1008
8.9006 | -104.1013
8.9043 | -104.1048
8.9081 | -104.112
8.9118 | -104.1167
8.9156 | -104.1177
8.9193 | -104.1214
8.9231 | -104.1258
8.9268 | -104.1318
8.9306 | -104.1401
8.9343 | -104.1467
8.9381 | -104.1498
8.9418 | -104.1521
8.9456 | -104.1542
8.9493 | -104.1518
8.9531 | -104.1463
8.9568 | -104.1432
8.9606 | -104.1385
8.9643 | -104.13
8.9681 | -104.1192
8.9718 | -104.1061
8.9756 | -104.0893
8.9793 | -104.0725
8.9831 | -104.0504
8.9868 | -104.0281
8.9906 | -104.0063
8.9943 | -103.9838
8.9981 | -103.9657
9.0018 | -103.9507
9.0056 | -103.9402
9.0093 | -103.9284
9.0131 | -103.9246
9.0168 | -103.9293
9.0206 | -103.9398
9.0243 | -103.9562
9.0281 | -103.9773
9.0318 | -104.0038
9.0356 | -104.0357
9.0393 | -104.0701
9.0431 | -104.102
9.0468 | -104.1362
9.0506 | -104.168
9.0543 | -104.1998
9.0581 | -104.2294
9.0618 | -104.2534
9.0656 | -104.2661
9.0693 | -104.2752
9.0731 | -104.2812
9.0768 | -104.2828
9.0806 | -104.2784
9.0843 | -104.2662
9.0881 | -104.2461
9.0918 | -104.218
9.0956 | -104.1912
9.0993 | -104.1613
9.1031 | -104.1321
9.1068 | -104.1078
9.1106 | -104.0765
9.1144 | -104.0542
9.1181 | -104.0371
9.1219 | -104.0196
9.1256 | -104.0073
9.1294 | -103.9995
9.1331 | -103.9952
9.1369 | -103.9946
9.1406 | -104.0013
9.1444 | -104.0064
9.1481 | -104.0137
9.1519 | -104.0243
9.1556 | -104.0362
9.1594 | -104.0543
9.1631 | -104.0719
9.1669 | -104.086
9.1706 | -104.0998
9.1744 | -104.113
9.1781 | -104.1234
9.1819 | -104.1291
9.1856 | -104.1281
9.1894 | -104.127
9.1931 | -104.1257
9.1969 | -104.1204
9.2006 | -104.1209
9.2044 | -104.1179
9.2081 | -104.1137
9.2119 | -104.1074
9.2156 | -104.0968
9.2194 | -104.0853
9.2231 | -104.075
9.2269 | -104.0665
9.2306 | -104.0566
9.2344 | -104.047
9.2381 | -104.0392
9.2419 | -104.0292
9.2456 | -104.0261
9.2494 | -104.0216
9.2531 | -104.0204
9.2569 | -104.0286
9.2606 | -104.0383
9.2644 | -104.0456
9.2681 | -104.0582
9.2719 | -104.0747
9.2756 | -104.0966
9.2794 | -104.1181
9.2831 | -104.1388
9.2869 | -104.1581
9.2906 | -104.179
9.2944 | -104.1982
9.2981 | -104.2216
9.3019 | -104.2429
9.3056 | -104.2639
9.3094 | -104.2868
9.3131 | -104.3097
9.3169 | -104.3317
9.3206 | -104.3533
9.3244 | -104.3723
9.3281 | -104.3913
9.3319 | -104.4063
9.3356 | -104.4167
9.3394 | -104.4218
9.3431 | -104.4259
9.3469 | -104.4292
9.3506 | -104.4286
9.3544 | -104.4253
9.3582 | -104.4187
9.3619 | -104.4056
9.3657 | -104.3893
9.3694 | -104.3704
9.3732 | -104.3523
9.3769 | -104.3285
9.3807 | -104.3061
9.3844 | -104.2809
9.3882 | -104.2542
9.3919 | -104.2313
9.3957 | -104.2091
9.3994 | -104.1901
9.4032 | -104.1795
9.4069 | -104.1683
9.4107 | -104.1575
9.4144 | -104.1506
9.4182 | -104.1522
9.4219 | -104.1564
9.4257 | -104.1628
9.4294 | -104.1726
9.4332 | -104.1862
9.4369 | -104.2009
9.4407 | -104.2182
9.4444 | -104.2368
9.4482 | -104.2578
9.4519 | -104.2794
9.4557 | -104.3043
9.4594 | -104.3296
9.4632 | -104.355
9.4669 | -104.3757
9.4707 | -104.3964
9.4744 | -104.415
9.4782 | -104.4287
9.4819 | -104.4385
9.4857 | -104.4432
9.4894 | -104.4473
9.4932 | -104.4475
9.4969 | -104.4388
9.5007 | -104.4282
9.5044 | -104.418
9.5082 | -104.4036
9.5119 | -104.3888
9.5157 | -104.3713
9.5194 | -104.3481
9.5232 | -104.3257
9.5269 | -104.3041
9.5307 | -104.2835
9.5344 | -104.2587
9.5382 | -104.234
9.5419 | -104.2125
9.5457 | -104.1968
9.5494 | -104.1861
9.5532 | -104.1738
9.5569 | -104.1657
9.5607 | -104.1638
9.5644 | -104.1662
9.5682 | -104.1751
9.5719 | -104.1871
9.5757 | -104.1993
9.5794 | -104.2117
9.5832 | -104.2255
9.5869 | -104.2434
9.5907 | -104.259
9.5944 | -104.2747
9.5982 | -104.289
9.602 | -104.3017
9.6057 | -104.3178
9.6095 | -104.332
9.6132 | -104.3447
9.617 | -104.361
9.6207 | -104.3728
9.6245 | -104.3841
9.6282 | -104.4004
9.632 | -104.4197
9.6357 | -104.4397
9.6395 | -104.4601
9.6432 | -104.479
9.647 | -104.497
9.6507 | -104.512
9.6545 | -104.5302
9.6582 | -104.551
9.662 | -104.5727
9.6657 | -104.5888
9.6695 | -104.6007
9.6732 | -104.6162
9.677 | -104.6349
9.6807 | -104.652
9.6845 | -104.6644
9.6882 | -104.6746
9.692 | -104.6879
9.6957 | -104.6984
9.6995 | -104.7051
9.7032 | -104.7105
9.707 | -104.7087
9.7107 | -104.7052
9.7145 | -104.7059
9.7182 | -104.7063
9.722 | -104.6997
9.7257 | -104.6865
9.7295 | -104.6683
9.7332 | -104.6466
9.737 | -104.6147
9.7407 | -104.5918
9.7445 | -104.5687
9.7482 | -104.5442
9.752 | -104.5194
9.7557 | -104.4979
9.7595 | -104.4756
9.7632 | -104.4554
9.767 | -104.4374
9.7707 | -104.4202
9.7745 | -104.407
9.7782 | -104.401
9.782 | -104.395
9.7857 | -104.3931
9.7895 | -104.3984
9.7932 | -104.4084
9.797 | -104.42
9.8007 | -104.4337
9.8045 | -104.4479
9.8082 | -104.4643
9.812 | -104.4876
9.8157 | -104.5119
9.8195 | -104.5347
9.8232 | -104.5582
9.827 | -104.578
9.8307 | -104.6012
9.8345 | -104.6227
9.8382 | -104.6405
9.842 | -104.6574
9.8458 | -104.6753
9.8495 | -104.6915
9.8533 | -104.7052
9.857 | -104.7133
9.8608 | -104.7168
9.8645 | -104.7211
9.8683 | -104.7236
9.872 | -104.728
9.8758 | -104.7272
9.8795 | -104.7231
9.8833 | -104.719
9.887 | -104.7113
9.8908 | -104.7044
9.8945 | -104.6978
9.8983 | -104.6909
9.902 | -104.6853
9.9058 | -104.6787
9.9095 | -104.6714
9.9133 | -104.6638
9.917 | -104.6601
9.9208 | -104.6572
9.9245 | -104.656
9.9283 | -104.6545
9.932 | -104.6522
9.9358 | -104.6494
9.9395 | -104.6489
9.9433 | -104.6491
9.947 | -104.6494
9.9508 | -104.6508
9.9545 | -104.6558
9.9583 | -104.658
9.962 | -104.6619
9.9658 | -104.6649
9.9695 | -104.6672
9.9733 | -104.6759
9.977 | -104.6761
9.9808 | -104.6741
9.9845 | -104.6723
9.9883 | -104.6719
9.992 | -104.6692
9.9958 | -104.6682
9.9995 | -104.6712
10.0033 | -104.6758
10.007 | -104.6824
10.0108 | -104.6886
10.0145 | -104.6983
10.0183 | -104.7084
10.022 | -104.724
10.0258 | -104.7359
10.0295 | -104.7415
10.0333 | -104.7476
10.037 | -104.7518
10.0408 | -104.7566
10.0445 | -104.7591
10.0483 | -104.7594
10.052 | -104.7587
10.0558 | -104.7581
10.0595 | -104.7593
10.0633 | -104.7567
10.067 | -104.7579
10.0708 | -104.7621
10.0745 | -104.769
10.0783 | -104.7769
10.082 | -104.7868
10.0858 | -104.7988
10.0895 | -104.8154
10.0933 | -104.8342
10.0971 | -104.8567
10.1008 | -104.8807
10.1046 | -104.9052
10.1083 | -104.9262
10.1121 | -104.9494
10.1158 | -104.9746
10.1196 | -105.0025
10.1233 | -105.0317
10.1271 | -105.057
10.1308 | -105.076
10.1346 | -105.0848
10.1383 | -105.0894
10.1421 | -105.0928
10.1458 | -105.0882
10.1496 | -105.0761
10.1533 | -105.0593
10.1571 | -105.0367
10.1608 | -105.0085
10.1646 | -104.974
10.1683 | -104.9339
10.1721 | -104.8957
10.1758 | -104.8593
10.1796 | -104.8187
10.1833 | -104.7736
10.1871 | -104.729
10.1908 | -104.6912
10.1946 | -104.6488
10.1983 | -104.6114
10.2021 | -104.5792
10.2058 | -104.5488
10.2096 | -104.5258
10.2133 | -104.5058
10.2171 | -104.4902
10.2208 | -104.4764
10.2246 | -104.4681
10.2283 | -104.4688
10.2321 | -104.471
10.2358 | -104.4793
10.2396 | -104.4926
10.2433 | -104.5076
10.2471 | -104.5247
10.2508 | -104.5469
10.2546 | -104.5765
10.2583 | -104.6047
10.2621 | -104.6329
10.2658 | -104.6593
10.2696 | -104.6872
10.2733 | -104.7149
10.2771 | -104.7427
10.2808 | -104.759
10.2846 | -104.7801
10.2883 | -104.7976
10.2921 | -104.8109
10.2958 | -104.8234
10.2996 | -104.8313
10.3033 | -104.8392
10.3071 | -104.8465
10.3108 | -104.8479
10.3146 | -104.8481
10.3183 | -104.8462
10.3221 | -104.8435
10.3258 | -104.838
10.3296 | -104.8315
10.3333 | -104.8261
10.3371 | -104.8225
10.3408 | -104.8192
10.3446 | -104.8213
10.3484 | -104.8197
10.3521 | -104.8162
10.3559 | -104.8152
10.3596 | -104.8147
10.3634 | -104.816
10.3671 | -104.8256
10.3709 | -104.8397
10.3746 | -104.8542
10.3784 | -104.8644
10.3821 | -104.8763
10.3859 | -104.8868
10.3896 | -104.8983
10.3934 | -104.9129
10.3971 | -104.9235
10.4009 | -104.9306
10.4046 | -104.9377
10.4084 | -104.943
10.4121 | -104.9434
10.4159 | -104.9407
10.4196 | -104.9366
10.4234 | -104.928
10.4271 | -104.9154
10.4309 | -104.9008
10.4346 | -104.8856
10.4384 | -104.8739
10.4421 | -104.8606
10.4459 | -104.849
10.4496 | -104.8394
10.4534 | -104.8351
10.4571 | -104.8257
10.4609 | -104.8162
10.4646 | -104.8119
10.4684 | -104.8112
10.4721 | -104.8119
10.4759 | -104.808
10.4796 | -104.8111
10.4834 | -104.8134
10.4871 | -104.8136
10.4909 | -104.8136
10.4946 | -104.8122
10.4984 | -104.8154
10.5021 | -104.821
10.5059 | -104.8291
10.5096 | -104.8378
10.5134 | -104.8424
10.5171 | -104.8459
10.5209 | -104.8416
10.5246 | -104.8393
10.5284 | -104.8375
10.5321 | -104.8326
10.5359 | -104.8254
10.5396 | -104.8163
10.5434 | -104.8036
10.5471 | -104.7871
10.5509 | -104.7731
10.5546 | -104.7574
10.5584 | -104.7354
10.5621 | -104.7132
10.5659 | -104.6927
10.5696 | -104.671
10.5734 | -104.6492
10.5771 | -104.6322
10.5809 | -104.6128
10.5846 | -104.592
10.5884 | -104.5717
10.5922 | -104.5558
10.5959 | -104.5431
10.5997 | -104.5323
10.6034 | -104.525
10.6072 | -104.5185
10.6109 | -104.5195
10.6147 | -104.5264
10.6184 | -104.5351
10.6222 | -104.5478
10.6259 | -104.5644
10.6297 | -104.5853
10.6334 | -104.599
10.6372 | -104.6157
10.6409 | -104.6362
10.6447 | -104.6606
10.6484 | -104.6875
10.6522 | -104.7116
10.6559 | -104.7346
10.6597 | -104.7546
10.6634 | -104.7773
10.6672 | -104.8016
10.6709 | -104.8177
10.6747 | -104.8321
10.6784 | -104.8436
10.6822 | -104.847
10.6859 | -104.8492
10.6897 | -104.8505
10.6934 | -104.8444
10.6972 | -104.8315
10.7009 | -104.8143
10.7047 | -104.7949
10.7084 | -104.7716
10.7122 | -104.7479
10.7159 | -104.7211
10.7197 | -104.6908
10.7234 | -104.657
10.7272 | -104.617
10.7309 | -104.5772
10.7347 | -104.5411
10.7384 | -104.5061
10.7422 | -104.4722
10.7459 | -104.4377
10.7497 | -104.4061
10.7534 | -104.3791
10.7572 | -104.3585
10.7609 | -104.3455
10.7647 | -104.3364
10.7684 | -104.3302
10.7722 | -104.3332
10.7759 | -104.341
10.7797 | -104.3551
10.7834 | -104.3801
10.7872 | -104.4123
10.7909 | -104.4441
10.7947 | -104.4712
10.7984 | -104.495
10.8022 | -104.5174
10.8059 | -104.5404
10.8097 | -104.5624
10.8134 | -104.5828
10.8172 | -104.5952
10.8209 | -104.5889
10.8247 | -104.5932
10.8285 | -104.5984
10.8322 | -104.5985
10.836 | -104.594
10.8397 | -104.5888
10.8435 | -104.582
10.8472 | -104.5727
10.851 | -104.5608
10.8547 | -104.5494
10.8585 | -104.5378
10.8622 | -104.527
10.866 | -104.5186
10.8697 | -104.512
10.8735 | -104.5019
10.8772 | -104.491
10.881 | -104.4892
10.8847 | -104.4916
10.8885 | -104.4914
10.8922 | -104.4901
10.896 | -104.4906
10.8997 | -104.4895
10.9035 | -104.4935
10.9072 | -104.4989
10.911 | -104.5018
10.9147 | -104.5037
10.9185 | -104.5059
10.9222 | -104.5067
10.926 | -104.5111
10.9297 | -104.5153
10.9335 | -104.5197
10.9372 | -104.5272
10.941 | -104.5338
10.9447 | -104.5441
10.9485 | -104.5616
10.9522 | -104.5851
10.956 | -104.6114
10.9597 | -104.6383
10.9635 | -104.6643
10.9672 | -104.6931
10.971 | -104.721
10.9747 | -104.7515
10.9785 | -104.783
10.9822 | -104.8128
10.986 | -104.8427
10.9897 | -104.8724
10.9935 | -104.8972
10.9972 | -104.9166
11.001 | -104.9331
11.0047 | -104.9458
11.0085 | -104.9537
11.0122 | -104.9607
11.016 | -104.9631
11.0197 | -104.9566
11.0235 | -104.9441
11.0272 | -104.9331
11.031 | -104.9194
11.0347 | -104.9064
11.0385 | -104.8928
11.0422 | -104.8742
11.046 | -104.8536
11.0497 | -104.8347
11.0535 | -104.8197
11.0572 | -104.804
11.061 | -104.7892
11.0647 | -104.7733
11.0685 | -104.7611
11.0722 | -104.7556
11.076 | -104.7504
11.0798 | -104.7452
11.0835 | -104.7435
11.0873 | -104.7389
11.091 | -104.736
11.0948 | -104.7367
11.0985 | -104.732
11.1023 | -104.7289
11.106 | -104.7352
11.1098 | -104.7426
11.1135 | -104.7491
11.1173 | -104.7552
11.121 | -104.7667
11.1248 | -104.7789
11.1285 | -104.7858
11.1323 | -104.7937
11.136 | -104.7999
11.1398 | -104.8052
11.1435 | -104.8085
11.1473 | -104.8104
11.151 | -104.8084
11.1548 | -104.8095
11.1585 | -104.8105
11.1623 | -104.8092
11.166 | -104.8073
11.1698 | -104.8085
11.1735 | -104.8098
11.1773 | -104.8116
11.181 | -104.817
11.1848 | -104.8248
11.1885 | -104.8297
11.1923 | -104.834
11.196 | -104.8441
11.1998 | -104.8568
11.2035 | -104.8663
11.2073 | -104.8796
11.211 | -104.8956
11.2148 | -104.9125
11.2185 | -104.9293
11.2223 | -104.9451
11.226 | -104.9617
11.2298 | -104.979
11.2335 | -104.9965
11.2373 | -105.0189
11.241 | -105.04
11.2448 | -105.0598
11.2485 | -105.079
11.2523 | -105.0961
11.256 | -105.1127
11.2598 | -105.1307
11.2635 | -105.1471
11.2673 | -105.1508
11.271 | -105.1558
11.2748 | -105.1639
11.2785 | -105.1677
11.2823 | -105.1679
11.286 | -105.1683
11.2898 | -105.1686
11.2935 | -105.1653
11.2973 | -105.159
11.301 | -105.1524
11.3048 | -105.1465
11.3085 | -105.1364
11.3123 | -105.1282
11.316 | -105.1194
11.3198 | -105.1118
11.3235 | -105.1057
11.3273 | -105.1014
11.3311 | -105.094
11.3348 | -105.0895
11.3386 | -105.0898
11.3423 | -105.0914
11.3461 | -105.0924
11.3498 | -105.0987
11.3536 | -105.1069
11.3573 | -105.1163
11.3611 | -105.1356
11.3648 | -105.148
11.3686 | -105.165
11.3723 | -105.1839
11.3761 | -105.1998
11.3798 | -105.2122
11.3836 | -105.229
11.3873 | -105.2473
11.3911 | -105.2662
11.3948 | -105.2826
11.3986 | -105.2936
11.4023 | -105.3038
11.4061 | -105.3166
11.4098 | -105.3272
11.4136 | -105.3345
11.4173 | -105.3392
11.4211 | -105.3422
11.4248 | -105.3435
11.4286 | -105.3405
11.4323 | -105.3388
11.4361 | -105.3304
11.4398 | -105.3215
11.4436 | -105.3142
11.4473 | -105.3065
11.4511 | -105.2978
11.4548 | -105.2908
11.4586 | -105.2849
11.4623 | -105.2796
11.4661 | -105.2727
11.4698 | -105.2681
11.4736 | -105.2637
11.4773 | -105.2589
11.4811 | -105.2539
11.4848 | -105.247
11.4886 | -105.2404
11.4923 | -105.2357
11.4961 | -105.2326
11.4998 | -105.2288
11.5036 | -105.2257
11.5073 | -105.2245
11.5111 | -105.2236
11.5148 | -105.2155
11.5186 | -105.2103
11.5223 | -105.2042
11.5261 | -105.1971
11.5298 | -105.1927
11.5336 | -105.1849
11.5373 | -105.1734
11.5411 | -105.1604
11.5448 | -105.1488
11.5486 | -105.1401
11.5523 | -105.1254
11.5561 | -105.1117
11.5598 | -105.0939
11.5636 | -105.0762
11.5673 | -105.0646
11.5711 | -105.0499
11.5749 | -105.0316
11.5786 | -105.0118
11.5824 | -104.9967
11.5861 | -104.9836
11.5899 | -104.9699
11.5936 | -104.9568
11.5974 | -104.948
11.6011 | -104.9413
11.6049 | -104.9387
11.6086 | -104.9346
11.6124 | -104.9293
11.6161 | -104.9244
11.6199 | -104.9214
11.6236 | -104.9185
11.6274 | -104.9159
11.6311 | -104.9149
11.6349 | -104.9113
11.6386 | -104.9024
11.6424 | -104.901
11.6461 | -104.897
11.6499 | -104.8925
11.6536 | -104.8917
11.6574 | -104.89
11.6611 | -104.887
11.6649 | -104.8825
11.6686 | -104.8794
11.6724 | -104.8732
11.6761 | -104.8647
11.6799 | -104.8534
11.6836 | -104.8472
11.6874 | -104.8413
11.6911 | -104.8354
11.6949 | -104.8265
11.6986 | -104.815
11.7024 | -104.8088
11.7061 | -104.8047
11.7099 | -104.8007
11.7136 | -104.7981
11.7174 | -104.7975
11.7211 | -104.8019
11.7249 | -104.8086
11.7286 | -104.8187
11.7324 | -104.8294
11.7361 | -104.8393
11.7399 | -104.8545
11.7436 | -104.8758
11.7474 | -104.9001
11.7511 | -104.9244
11.7549 | -104.9444
11.7586 | -104.9638
11.7624 | -104.9857
11.7661 | -105.0053
11.7699 | -105.0205
11.7736 | -105.0347
11.7774 | -105.0478
11.7811 | -105.0598
11.7849 | -105.067
11.7886 | -105.0698
11.7924 | -105.0724
11.7961 | -105.0755
11.7999 | -105.0759
11.8036 | -105.076
11.8074 | -105.0796
11.8111 | -105.0727
11.8149 | -105.0658
11.8187 | -105.0629
11.8224 | -105.0594
11.8262 | -105.0561
11.8299 | -105.0538
11.8337 | -105.0583
11.8374 | -105.0604
11.8412 | -105.0653
11.8449 | -105.0725
11.8487 | -105.0826
11.8524 | -105.0937
11.8562 | -105.1012
11.8599 | -105.1113
11.8637 | -105.1248
11.8674 | -105.142
11.8712 | -105.1537
11.8749 | -105.1663
11.8787 | -105.1753
11.8824 | -105.1801
11.8862 | -105.1842
11.8899 | -105.1881
11.8937 | -105.1927
11.8974 | -105.1956
11.9012 | -105.1906
11.9049 | -105.1944
11.9087 | -105.1938
11.9124 | -105.1896
11.9162 | -105.1849
11.9199 | -105.1771
11.9237 | -105.1708
11.9274 | -105.1704
11.9312 | -105.1718
11.9349 | -105.17
11.9387 | -105.1663
11.9424 | -105.1658
11.9462 | -105.1629
11.9499 | -105.1618
11.9537 | -105.1619
11.9574 | -105.1625
11.9612 | -105.1629
11.9649 | -105.1623
11.9687 | -105.1602
11.9724 | -105.1583
11.9762 | -105.1562
11.9799 | -105.1503
11.9837 | -105.1468
11.9874 | -105.1419
11.9912 | -105.1289
11.9949 | -105.1143
11.9987 | -105.1004
12.0024 | -105.0833
12.0062 | -105.0625
12.0099 | -105.0452
12.0137 | -105.0301
12.0174 | -105.0167
12.0212 | -105.0069
12.0249 | -104.998
12.0287 | -104.9924
12.0324 | -104.9879
12.0362 | -104.9897
12.0399 | -104.9928
12.0437 | -104.9979
12.0474 | -105.0055
12.0512 | -105.0197
12.0549 | -105.0357
12.0587 | -105.0499
12.0625 | -105.0629
12.0662 | -105.0717
12.07 | -105.0834
12.0737 | -105.0992
12.0775 | -105.1154
12.0812 | -105.1268
12.085 | -105.1412
12.0887 | -105.1513
12.0925 | -105.1544
12.0962 | -105.1562
12.1 | -105.1528
12.1037 | -105.1461
12.1075 | -105.1349
12.1112 | -105.1215
12.115 | -105.1092
12.1187 | -105.0956
12.1225 | -105.08
12.1262 | -105.0614
12.13 | -105.0411
12.1337 | -105.0193
12.1375 | -104.9949
12.1412 | -104.9707
12.145 | -104.948
12.1487 | -104.9277
12.1525 | -104.9122
12.1562 | -104.8983
12.16 | -104.8835
12.1637 | -104.8729
12.1675 | -104.8665
12.1712 | -104.8601
12.175 | -104.8547
12.1787 | -104.8536
12.1825 | -104.8539
12.1862 | -104.8531
12.19 | -104.8495
12.1937 | -104.8476
12.1975 | -104.8463
12.2012 | -104.8465
12.205 | -104.8494
12.2087 | -104.8538
12.2125 | -104.8581
12.2162 | -104.8676
12.22 | -104.8811
12.2237 | -104.8937
12.2275 | -104.9086
12.2312 | -104.9289
12.235 | -104.9505
12.2387 | -104.9716
12.2425 | -104.9962
12.2462 | -105.0208
12.25 | -105.0452
12.2537 | -105.0695
12.2575 | -105.0962
12.2612 | -105.1203
12.265 | -105.1488
12.2687 | -105.1768
12.2725 | -105.2019
12.2762 | -105.2247
12.28 | -105.2436
12.2837 | -105.2603
12.2875 | -105.2746
12.2912 | -105.2848
12.295 | -105.2937
12.2987 | -105.3022
12.3025 | -105.311
12.3063 | -105.3189
12.31 | -105.3272
12.3138 | -105.3364
12.3175 | -105.3418
12.3213 | -105.3469
12.325 | -105.3521
12.3288 | -105.3543
12.3325 | -105.3581
12.3363 | -105.3658
12.34 | -105.3726
12.3438 | -105.3804
12.3475 | -105.393
12.3513 | -105.4092
12.355 | -105.4227
12.3588 | -105.4346
12.3625 | -105.4437
12.3663 | -105.4558
12.37 | -105.4655
12.3738 | -105.474
12.3775 | -105.478
12.3813 | -105.4799
12.385 | -105.4768
12.3888 | -105.4767
12.3925 | -105.4746
12.3963 | -105.469
12.4 | -105.463
12.4038 | -105.4507
12.4075 | -105.4305
12.4113 | -105.4084
12.415 | -105.3853
12.4188 | -105.3571
12.4225 | -105.3285
12.4263 | -105.2966
12.43 | -105.2607
12.4338 | -105.2246
12.4375 | -105.189
12.4413 | -105.1484
12.445 | -105.1123
12.4488 | -105.0779
12.4525 | -105.0443
12.4563 | -105.0085
12.46 | -104.9762
12.4638 | -104.9448
12.4675 | -104.9183
12.4713 | -104.8918
12.475 | -104.8631
12.4788 | -104.8413
12.4825 | -104.82
12.4863 | -104.8005
12.49 | -104.7852
12.4938 | -104.7727
12.4975 | -104.7633
12.5013 | -104.7557
12.505 | -104.7508
12.5088 | -104.7487
12.5125 | -104.7505
12.5163 | -104.7518
12.52 | -104.7525
12.5238 | -104.7585
12.5275 | -104.7632
12.5313 | -104.7676
12.535 | -104.7764
12.5388 | -104.7924
12.5425 | -104.8099
12.5463 | -104.8269
12.55 | -104.8455
12.5538 | -104.8661
12.5576 | -104.8895
12.5613 | -104.9154
12.5651 | -104.9437
12.5688 | -104.9683
12.5726 | -104.9919
12.5763 | -105.0165
12.5801 | -105.0439
12.5838 | -105.0693
12.5876 | -105.0939
12.5913 | -105.1188
12.5951 | -105.1444
12.5988 | -105.1682
12.6026 | -105.1872
12.6063 | -105.2057
12.6101 | -105.2263
12.6138 | -105.2482
12.6176 | -105.2688
12.6213 | -105.2888
12.6251 | -105.3042
12.6288 | -105.3191
12.6326 | -105.3341
12.6363 | -105.3501
12.6401 | -105.3672
12.6438 | -105.3816
12.6476 | -105.3957
12.6513 | -105.4089
12.6551 | -105.4216
12.6588 | -105.4348
12.6626 | -105.4405
12.6663 | -105.4442
12.6701 | -105.4493
12.6738 | -105.4505
12.6776 | -105.4467
12.6813 | -105.4421
12.6851 | -105.4392
12.6888 | -105.4344
12.6926 | -105.428
12.6963 | -105.4149
12.7001 | -105.3991
12.7038 | -105.3836
12.7076 | -105.3663
12.7113 | -105.3504
12.7151 | -105.3321
12.7188 | -105.3124
12.7226 | -105.2885
12.7263 | -105.2657
12.7301 | -105.2427
12.7338 | -105.2175
12.7376 | -105.1923
12.7413 | -105.1691
12.7451 | -105.1463
12.7488 | -105.1273
12.7526 | -105.1098
12.7563 | -105.097
12.7601 | -105.0897
12.7638 | -105.084
12.7676 | -105.0807
12.7713 | -105.0809
12.7751 | -105.0819
12.7788 | -105.0868
12.7826 | -105.0963
12.7863 | -105.1049
12.7901 | -105.1172
12.7938 | -105.1326
12.7976 | -105.1449
12.8013 | -105.1589
12.8051 | -105.1767
12.8089 | -105.1954
12.8126 | -105.2148
12.8164 | -105.2323
12.8201 | -105.2483
12.8239 | -105.2621
12.8276 | -105.2776
12.8314 | -105.2941
12.8351 | -105.3061
12.8389 | -105.3181
12.8426 | -105.3293
12.8464 | -105.3406
12.8501 | -105.3529
12.8539 | -105.3608
12.8576 | -105.3672
12.8614 | -105.3793
12.8651 | -105.39
12.8689 | -105.3943
12.8726 | -105.3975
12.8764 | -105.3999
12.8801 | -105.4039
12.8839 | -105.4067
12.8876 | -105.4105
12.8914 | -105.4117
12.8951 | -105.4107
12.8989 | -105.4114
12.9026 | -105.4116
12.9064 | -105.4115
12.9101 | -105.4133
12.9139 | -105.4134
12.9176 | -105.4132
12.9214 | -105.4128
12.9251 | -105.4106
12.9289 | -105.406
12.9326 | -105.4008
12.9364 | -105.3948
12.9401 | -105.3869
12.9439 | -105.3785
12.9476 | -105.368
12.9514 | -105.3601
12.9551 | -105.353
12.9589 | -105.3461
12.9626 | -105.3373
12.9664 | -105.3309
12.9701 | -105.3267
12.9739 | -105.3185
12.9776 | -105.3121
12.9814 | -105.31
12.9851 | -105.3169
12.9889 | -105.318
12.9926 | -105.3159
12.9964 | -105.3167
13.0001 | -105.3193
13.0039 | -105.3214
13.0076 | -105.3266
13.0114 | -105.3363
13.0151 | -105.3444
13.0189 | -105.351
13.0226 | -105.3594
13.0264 | -105.3697
13.0301 | -105.3804
13.0339 | -105.391
13.0376 | -105.4004
13.0414 | -105.411
13.0451 | -105.4222
13.0489 | -105.4273
13.0527 | -105.4367
13.0564 | -105.4502
13.0602 | -105.4645
13.0639 | -105.4769
13.0677 | -105.491
13.0714 | -105.5048
13.0752 | -105.5152
13.0789 | -105.5235
13.0827 | -105.5336
13.0864 | -105.5428
13.0902 | -105.5526
13.0939 | -105.56
13.0977 | -105.5645
13.1014 | -105.5703
13.1052 | -105.5754
13.1089 | -105.5777
13.1127 | -105.5813
13.1164 | -105.5828
13.1202 | -105.5812
13.1239 | -105.5837
13.1277 | -105.5858
13.1314 | -105.5847
13.1352 | -105.5826
13.1389 | -105.5806
13.1427 | -105.5744
13.1464 | -105.5687
13.1502 | -105.5648
13.1539 | -105.56
13.1577 | -105.5552
13.1614 | -105.552
13.1652 | -105.5503
13.1689 | -105.5477
13.1727 | -105.5443
13.1764 | -105.5383
13.1802 | -105.534
13.1839 | -105.5319
13.1877 | -105.527
13.1914 | -105.5165
13.1952 | -105.5081
13.1989 | -105.4991
13.2027 | -105.4854
13.2064 | -105.4729
13.2102 | -105.4622
13.2139 | -105.4533
13.2177 | -105.4424
13.2214 | -105.4306
13.2252 | -105.4167
13.2289 | -105.4035
13.2327 | -105.3899
13.2364 | -105.3763
13.2402 | -105.3605
13.2439 | -105.3378
13.2477 | -105.3147
13.2514 | -105.2915
13.2552 | -105.2697
13.2589 | -105.2456
13.2627 | -105.2208
13.2664 | -105.1944
13.2702 | -105.1708
13.2739 | -105.144
13.2777 | -105.1158
13.2814 | -105.0855
13.2852 | -105.0563
13.2889 | -105.0283
13.2927 | -104.998
13.2965 | -104.9693
13.3002 | -104.9424
13.304 | -104.9142
13.3077 | -104.8897
13.3115 | -104.8683
13.3152 | -104.847
13.319 | -104.8292
13.3227 | -104.8138
13.3265 | -104.8018
13.3302 | -104.7903
13.334 | -104.7844
13.3377 | -104.7799
13.3415 | -104.7782
13.3452 | -104.7794
13.349 | -104.7835
13.3527 | -104.7909
13.3565 | -104.7992
13.3602 | -104.8116
13.364 | -104.8247
13.3677 | -104.843
13.3715 | -104.8658
13.3752 | -104.887
13.379 | -104.9121
13.3827 | -104.9428
13.3865 | -104.9717
13.3902 | -104.9994
13.394 | -105.0311
13.3977 | -105.0638
13.4015 | -105.0946
13.4052 | -105.1236
13.409 | -105.1539
13.4127 | -105.1832
13.4165 | -105.2108
13.4202 | -105.2383
13.424 | -105.2642
13.4277 | -105.2864
13.4315 | -105.3105
13.4352 | -105.3333
13.439 | -105.3522
13.4427 | -105.3659
13.4465 | -105.3781
13.4502 | -105.3886
13.454 | -105.394
13.4577 | -105.3983
13.4615 | -105.4006
13.4652 | -105.4026
13.469 | -105.402
13.4727 | -105.4006
13.4765 | -105.3985
13.4802 | -105.3919
13.484 | -105.3867
13.4877 | -105.3798
13.4915 | -105.3714
13.4952 | -105.3628
13.499 | -105.3531
13.5027 | -105.3414
13.5065 | -105.3305
13.5102 | -105.3197
13.514 | -105.3094
13.5177 | -105.2992
13.5215 | -105.2914
13.5252 | -105.2834
13.529 | -105.2733
13.5327 | -105.2644
13.5365 | -105.2548
13.5403 | -105.2447
13.544 | -105.2342
13.5478 | -105.2262
13.5515 | -105.2184
13.5553 | -105.2084
13.559 | -105.2
13.5628 | -105.1901
13.5665 | -105.1826
13.5703 | -105.1763
13.574 | -105.1675
13.5778 | -105.1565
13.5815 | -105.1457
13.5853 | -105.1355
13.589 | -105.1237
13.5928 | -105.1092
13.5965 | -105.096
13.6003 | -105.084
13.604 | -105.0709
13.6078 | -105.0582
13.6115 | -105.0475
13.6153 | -105.0349
13.619 | -105.0215
13.6228 | -105.0072
13.6265 | -104.9931
13.6303 | -104.9794
13.634 | -104.9677
13.6378 | -104.959
13.6415 | -104.9504
13.6453 | -104.9395
13.649 | -104.9322
13.6528 | -104.926
13.6565 | -104.9181
13.6603 | -104.9161
13.664 | -104.9171
13.6678 | -104.9187
13.6715 | -104.9224
13.6753 | -104.931
13.679 | -104.939
13.6828 | -104.949
13.6865 | -104.9631
13.6903 | -104.9775
13.694 | -104.9906
13.6978 | -105.0051
13.7015 | -105.0181
13.7053 | -105.0306
13.709 | -105.0473
13.7128 | -105.064
13.7165 | -105.0772
13.7203 | -105.0898
13.724 | -105.1007
13.7278 | -105.1118
13.7315 | -105.123
13.7353 | -105.1331
13.739 | -105.145
13.7428 | -105.1562
13.7465 | -105.1599
13.7503 | -105.1642
13.754 | -105.1721
13.7578 | -105.18
13.7615 | -105.1857
13.7653 | -105.1856
13.769 | -105.1874
13.7728 | -105.1891
13.7765 | -105.191
13.7803 | -105.1955
13.7841 | -105.1995
13.7878 | -105.2023
13.7916 | -105.2061
13.7953 | -105.2101
13.7991 | -105.2138
13.8028 | -105.2182
13.8066 | -105.2243
13.8103 | -105.2298
13.8141 | -105.2386
13.8178 | -105.2482
13.8216 | -105.2616
13.8253 | -105.2749
13.8291 | -105.286
13.8328 | -105.296
13.8366 | -105.3046
13.8403 | -105.31
13.8441 | -105.3185
13.8478 | -105.3291
13.8516 | -105.3378
13.8553 | -105.3459
13.8591 | -105.3483
13.8628 | -105.3478
13.8666 | -105.3445
13.8703 | -105.3385
13.8741 | -105.3321
13.8778 | -105.3228
13.8816 | -105.3092
13.8853 | -105.2923
13.8891 | -105.276
13.8928 | -105.2601
13.8966 | -105.2403
13.9003 | -105.2201
13.9041 | -105.2003
13.9078 | -105.1753
13.9116 | -105.1502
13.9153 | -105.1293
13.9191 | -105.1105
13.9228 | -105.0928
13.9266 | -105.0728
13.9303 | -105.0513
13.9341 | -105.0335
13.9378 | -105.0177
13.9416 | -105.0052
13.9453 | -104.9975
13.9491 | -104.9907
13.9528 | -104.9851
13.9566 | -104.9779
13.9603 | -104.9708
13.9641 | -104.9642
13.9678 | -104.9627
13.9716 | -104.9639
13.9753 | -104.9664
13.9791 | -104.9688
13.9828 | -104.9708
13.9866 | -104.9733
13.9903 | -104.976
13.9941 | -104.9771
13.9978 | -104.976
14.0016 | -104.9763
14.0053 | -104.9753
14.0091 | -104.9728
14.0128 | -104.9644
14.0166 | -104.9559
14.0203 | -104.9471
14.0241 | -104.9364
14.0278 | -104.926
14.0316 | -104.9159
14.0354 | -104.9056
14.0391 | -104.8956
14.0429 | -104.8856
14.0466 | -104.8752
14.0504 | -104.8672
14.0541 | -104.8614
14.0579 | -104.8579
14.0616 | -104.8604
14.0654 | -104.8807
14.0691 | -104.8901
14.0729 | -104.9008
14.0766 | -104.9124
14.0804 | -104.931
14.0841 | -104.9526
14.0879 | -104.975
14.0916 | -104.9984
14.0954 | -105.0254
14.0991 | -105.0562
14.1029 | -105.0877
14.1066 | -105.1198
14.1104 | -105.1534
14.1141 | -105.1871
14.1179 | -105.2194
14.1216 | -105.2496
14.1254 | -105.2819
14.1291 | -105.3178
14.1329 | -105.3513
14.1366 | -105.3822
14.1404 | -105.4132
14.1441 | -105.4437
14.1479 | -105.4741
14.1516 | -105.503
14.1554 | -105.5302
14.1591 | -105.5598
14.1629 | -105.5892
14.1666 | -105.6187
14.1704 | -105.6411
14.1741 | -105.6592
14.1779 | -105.6755
14.1816 | -105.6897
14.1854 | -105.7043
14.1891 | -105.7174
14.1929 | -105.7301
14.1966 | -105.7422
14.2004 | -105.7493
14.2041 | -105.7528
14.2079 | -105.7605
14.2116 | -105.7664
14.2154 | -105.7726
14.2191 | -105.7816
14.2229 | -105.788
14.2266 | -105.7931
14.2304 | -105.7954
14.2341 | -105.7922
14.2379 | -105.7917
14.2416 | -105.7893
14.2454 | -105.7825
14.2491 | -105.7732
14.2529 | -105.7605
14.2566 | -105.7482
14.2604 | -105.7376
14.2641 | -105.7207
14.2679 | -105.701
14.2716 | -105.6796
14.2754 | -105.6573
14.2791 | -105.6327
14.2829 | -105.608
14.2867 | -105.5802
14.2904 | -105.5519
14.2942 | -105.523
14.2979 | -105.4955
14.3017 | -105.4632
14.3054 | -105.4311
14.3092 | -105.4029
14.3129 | -105.3739
14.3167 | -105.345
14.3204 | -105.3178
14.3242 | -105.2923
14.3279 | -105.2687
14.3317 | -105.2429
14.3354 | -105.2198
14.3392 | -105.1998
14.3429 | -105.1831
14.3467 | -105.1658
14.3504 | -105.1507
14.3542 | -105.1376
14.3579 | -105.127
14.3617 | -105.1177
14.3654 | -105.1139
14.3692 | -105.1121
14.3729 | -105.11
14.3767 | -105.112
14.3804 | -105.1158
14.3842 | -105.1199
14.3879 | -105.1261
14.3917 | -105.1344
14.3954 | -105.1445
14.3992 | -105.1547
14.4029 | -105.1674
14.4067 | -105.1846
14.4104 | -105.2026
14.4142 | -105.2175
14.4179 | -105.234
14.4217 | -105.2522
14.4254 | -105.2676
14.4292 | -105.2828
14.4329 | -105.2988
14.4367 | -105.3155
14.4404 | -105.3328
14.4442 | -105.3455
14.4479 | -105.3585
14.4517 | -105.3692
14.4554 | -105.3781
14.4592 | -105.3877
14.4629 | -105.3955
14.4667 | -105.4001
14.4704 | -105.4109
14.4742 | -105.4208
14.4779 | -105.4273
14.4817 | -105.4317
14.4854 | -105.4383
14.4892 | -105.441
14.4929 | -105.4427
14.4967 | -105.4452
14.5004 | -105.4475
14.5042 | -105.45
14.5079 | -105.4542
14.5117 | -105.4544
14.5154 | -105.4513
14.5192 | -105.4481
14.5229 | -105.4448
14.5267 | -105.4389
14.5305 | -105.4316
14.5342 | -105.4258
14.538 | -105.4216
14.5417 | -105.4152
14.5455 | -105.4117
14.5492 | -105.4036
14.553 | -105.3982
14.5567 | -105.393
14.5605 | -105.386
14.5642 | -105.3793
14.568 | -105.3744
14.5717 | -105.3659
14.5755 | -105.3584
14.5792 | -105.3553
14.583 | -105.3505
14.5867 | -105.3456
14.5905 | -105.3421
14.5942 | -105.3386
14.598 | -105.3343
14.6017 | -105.3334
14.6055 | -105.3326
14.6092 | -105.3437
14.613 | -105.3474
14.6167 | -105.3521
14.6205 | -105.3566
14.6242 | -105.361
14.628 | -105.3713
14.6317 | -105.3826
14.6355 | -105.3927
14.6392 | -105.4046
14.643 | -105.4187
14.6467 | -105.4335
14.6505 | -105.4481
14.6542 | -105.4608
14.658 | -105.4728
14.6617 | -105.4832
14.6655 | -105.4933
14.6692 | -105.5014
14.673 | -105.5084
14.6767 | -105.5151
14.6805 | -105.5171
14.6842 | -105.5182
14.688 | -105.5188
14.6917 | -105.515
14.6955 | -105.509
14.6992 | -105.5018
14.703 | -105.49
14.7067 | -105.4777
14.7105 | -105.4636
14.7142 | -105.4447
14.718 | -105.4286
14.7217 | -105.4144
14.7255 | -105.3991
14.7292 | -105.3847
14.733 | -105.37
14.7367 | -105.3554
14.7405 | -105.3397
14.7442 | -105.3232
14.748 | -105.3089
14.7517 | -105.2971
14.7555 | -105.2855
14.7592 | -105.2771
14.763 | -105.2692
14.7667 | -105.2638
14.7705 | -105.2634
14.7743 | -105.2645
14.778 | -105.2664
14.7818 | -105.2711
14.7855 | -105.2773
14.7893 | -105.283
14.793 | -105.2904
14.7968 | -105.3002
14.8005 | -105.3127
14.8043 | -105.3254
14.808 | -105.3398
14.8118 | -105.3517
14.8155 | -105.3623
14.8193 | -105.3745
14.823 | -105.3857
14.8268 | -105.3985
14.8305 | -105.4109
14.8343 | -105.4249
14.838 | -105.4388
14.8418 | -105.4481
14.8455 | -105.4573
14.8493 | -105.4683
14.853 | -105.4783
14.8568 | -105.485
14.8605 | -105.4929
14.8643 | -105.4993
14.868 | -105.5044
14.8718 | -105.5085
14.8755 | -105.5147
14.8793 | -105.5182
14.883 | -105.5194
14.8868 | -105.5185
14.8905 | -105.5174
14.8943 | -105.5167
14.898 | -105.5146
14.9018 | -105.5122
14.9055 | -105.5127
14.9093 | -105.5096
14.913 | -105.5066
14.9168 | -105.5028
14.9205 | -105.4972
14.9243 | -105.4919
14.928 | -105.4908
14.9318 | -105.4866
14.9355 | -105.481
14.9393 | -105.4743
14.943 | -105.468
14.9468 | -105.463
14.9505 | -105.4573
14.9543 | -105.45
14.958 | -105.4448
14.9618 | -105.4387
14.9655 | -105.431
14.9693 | -105.4237
14.973 | -105.4178
14.9768 | -105.4131
14.9805 | -105.4045
14.9843 | -105.3942
14.988 | -105.3823
14.9918 | -105.3744
14.9955 | -105.3671
14.9993 | -105.3625
15.003 | -105.3581
15.0068 | -105.3507
15.0105 | -105.3458
15.0143 | -105.3423
15.0181 | -105.3417
15.0218 | -105.3407
15.0256 | -105.3406
15.0293 | -105.3426
15.0331 | -105.3434
15.0368 | -105.3429
15.0406 | -105.3415
15.0443 | -105.343
15.0481 | -105.3451
15.0518 | -105.3462
15.0556 | -105.3487
15.0593 | -105.35
15.0631 | -105.3475
15.0668 | -105.3436
15.0706 | -105.3398
15.0743 | -105.3327
15.0781 | -105.3259
15.0818 | -105.3187
15.0856 | -105.308
15.0893 | -105.2936
15.0931 | -105.2781
15.0968 | -105.2616
15.1006 | -105.2466
15.1043 | -105.2302
15.1081 | -105.2117
15.1118 | -105.1962
15.1156 | -105.183
15.1193 | -105.1697
15.1231 | -105.1572
15.1268 | -105.1449
15.1306 | -105.1356
15.1343 | -105.1273
15.1381 | -105.1203
15.1418 | -105.1113
15.1456 | -105.102
15.1493 | -105.1042
15.1531 | -105.099
15.1568 | -105.0984
15.1606 | -105.0974
15.1643 | -105.0983
15.1681 | -105.1003
15.1718 | -105.1035
15.1756 | -105.1072
15.1793 | -105.1119
15.1831 | -105.1181
15.1868 | -105.125
15.1906 | -105.1318
15.1943 | -105.1396
15.1981 | -105.1513
15.2018 | -105.1616
15.2056 | -105.1739
15.2093 | -105.1834
15.2131 | -105.1926
15.2168 | -105.2002
15.2206 | -105.2082
15.2243 | -105.2185
15.2281 | -105.2236
15.2318 | -105.2282
15.2356 | -105.2283
15.2393 | -105.2273
15.2431 | -105.2277
15.2468 | -105.2269
15.2506 | -105.2252
15.2543 | -105.2255
15.2581 | -105.2256
15.2619 | -105.2234
15.2656 | -105.2196
15.2694 | -105.2178
15.2731 | -105.2156
15.2769 | -105.2145
15.2806 | -105.2143
15.2844 | -105.2125
15.2881 | -105.2109
15.2919 | -105.2104
15.2956 | -105.2098
15.2994 | -105.2119
15.3031 | -105.2145
15.3069 | -105.2175
15.3106 | -105.2198
15.3144 | -105.2259
15.3181 | -105.2316
15.3219 | -105.2361
15.3256 | -105.2418
15.3294 | -105.246
15.3331 | -105.2478
15.3369 | -105.2515
15.3406 | -105.2567
15.3444 | -105.2614
15.3481 | -105.2643
15.3519 | -105.2663
15.3556 | -105.266
15.3594 | -105.2676
15.3631 | -105.2709
15.3669 | -105.2714
15.3706 | -105.2711
15.3744 | -105.27
15.3781 | -105.2634
15.3819 | -105.2579
15.3856 | -105.2533
15.3894 | -105.2481
15.3931 | -105.2395
15.3969 | -105.2297
15.4006 | -105.2211
15.4044 | -105.2136
15.4081 | -105.2065
15.4119 | -105.201
15.4156 | -105.193
15.4194 | -105.1864
15.4231 | -105.1839
15.4269 | -105.1809
15.4306 | -105.1809
15.4344 | -105.1795
15.4381 | -105.1809
15.4419 | -105.1852
15.4456 | -105.1893
15.4494 | -105.1984
15.4531 | -105.2063
15.4569 | -105.2176
15.4606 | -105.2298
15.4644 | -105.2382
15.4681 | -105.2476
15.4719 | -105.2593
15.4756 | -105.2682
15.4794 | -105.276
15.4831 | -105.2857
15.4869 | -105.2939
15.4906 | -105.3002
15.4944 | -105.3044
15.4981 | -105.3099
15.5019 | -105.3184
15.5056 | -105.3258
15.5094 | -105.3292
15.5132 | -105.3279
15.5169 | -105.3246
15.5207 | -105.3166
15.5244 | -105.31
15.5282 | -105.3028
15.5319 | -105.2968
15.5357 | -105.2848
15.5394 | -105.2675
15.5432 | -105.2532
15.5469 | -105.2365
15.5507 | -105.2183
15.5544 | -105.2007
15.5582 | -105.1812
15.5619 | -105.1599
15.5657 | -105.139
15.5694 | -105.119
15.5732 | -105.0963
15.5769 | -105.0755
15.5807 | -105.0526
15.5844 | -105.0311
15.5882 | -105.0094
15.5919 | -104.9861
15.5957 | -104.9632
15.5994 | -104.9421
15.6032 | -104.9215
15.6069 | -104.9017
15.6107 | -104.8827
15.6144 | -104.8674
15.6182 | -104.8509
15.6219 | -104.8337
15.6257 | -104.8167
15.6294 | -104.7997
15.6332 | -104.7846
15.6369 | -104.7708
15.6407 | -104.7576
15.6444 | -104.7457
15.6482 | -104.7341
15.6519 | -104.7196
15.6557 | -104.7095
15.6594 | -104.6978
15.6632 | -104.6855
15.6669 | -104.6762
15.6707 | -104.6672
15.6744 | -104.6571
15.6782 | -104.6508
15.6819 | -104.6464
15.6857 | -104.6452
15.6894 | -104.6478
15.6932 | -104.6477
15.6969 | -104.6462
15.7007 | -104.6473
15.7044 | -104.6496
15.7082 | -104.6524
15.7119 | -104.6583
15.7157 | -104.6655
15.7194 | -104.6727
15.7232 | -104.6782
15.7269 | -104.6836
15.7307 | -104.6913
15.7344 | -104.7015
15.7382 | -104.7132
15.7419 | -104.7265
15.7457 | -104.7418
15.7494 | -104.7555
15.7532 | -104.7674
15.757 | -104.7786
15.7607 | -104.7923
15.7645 | -104.8092
15.7682 | -104.8256
15.772 | -104.8436
15.7757 | -104.8599
15.7795 | -104.8752
15.7832 | -104.8923
15.787 | -104.9086
15.7907 | -104.9242
15.7945 | -104.9411
15.7982 | -104.9595
15.802 | -104.9758
15.8057 | -104.9902
15.8095 | -105.0028
15.8132 | -105.0163
15.817 | -105.0299
15.8207 | -105.0415
15.8245 | -105.0527
15.8282 | -105.0603
15.832 | -105.0646
15.8357 | -105.0686
15.8395 | -105.0709
15.8432 | -105.0722
15.847 | -105.0707
15.8507 | -105.068
15.8545 | -105.0667
15.8582 | -105.063
15.862 | -105.0596
15.8657 | -105.0543
15.8695 | -105.045
15.8732 | -105.0359
15.877 | -105.0291
15.8807 | -105.0204
15.8845 | -105.0128
15.8882 | -105.0044
15.892 | -104.9967
15.8957 | -104.9879
15.8995 | -104.978
15.9032 | -104.9646
15.907 | -104.9525
15.9107 | -104.9407
15.9145 | -104.9303
15.9182 | -104.923
15.922 | -104.9142
15.9257 | -104.9052
15.9295 | -104.8973
15.9332 | -104.8937
15.937 | -104.8886
15.9407 | -104.8827
15.9445 | -104.8749
15.9482 | -104.8708
15.952 | -104.8683
15.9557 | -104.8681
15.9595 | -104.8691
15.9632 | -104.8697
15.967 | -104.8722
15.9707 | -104.8746
15.9745 | -104.8762
15.9782 | -104.8807
15.982 | -104.8886
15.9857 | -104.8969
15.9895 | -104.9057
15.9932 | -104.913
15.997 | -104.9161
16.0008 | -104.9159
16.0045 | -104.919
16.0083 | -104.9249
16.012 | -104.9256
16.0158 | -104.9266
16.0195 | -104.9275
16.0233 | -104.9205
16.027 | -104.9103
16.0308 | -104.9
16.0345 | -104.8886
16.0383 | -104.874
16.042 | -104.8595
16.0458 | -104.8417
16.0495 | -104.8211
16.0533 | -104.7962
16.057 | -104.7665
16.0608 | -104.7324
16.0645 | -104.6963
16.0683 | -104.6579
16.072 | -104.6169
16.0758 | -104.5714
16.0795 | -104.5269
16.0833 | -104.4815
16.087 | -104.4331
16.0908 | -104.3814
16.0945 | -104.3276
16.0983 | -104.2752
16.102 | -104.2269
16.1058 | -104.1785
16.1095 | -104.1285
16.1133 | -104.0785
16.117 | -104.0304
16.1208 | -103.9825
16.1245 | -103.9374
16.1283 | -103.8978
16.132 | -103.8598
16.1358 | -103.823
16.1395 | -103.7894
16.1433 | -103.7582
16.147 | -103.7294
16.1508 | -103.705
16.1545 | -103.6853
16.1583 | -103.6682
16.162 | -103.6523
16.1658 | -103.6414
16.1695 | -103.636
16.1733 | -103.6358
16.177 | -103.6378
16.1808 | -103.6416
16.1845 | -103.6488
16.1883 | -103.6574
16.192 | -103.6687
16.1958 | -103.6825
16.1995 | -103.6995
16.2033 | -103.7202
16.207 | -103.742
16.2108 | -103.7632
16.2145 | -103.7861
16.2183 | -103.8112
16.222 | -103.8382
16.2258 | -103.8699
16.2295 | -103.9037
16.2333 | -103.9315
16.237 | -103.9599
16.2408 | -103.9825
16.2446 | -104.006
16.2483 | -104.0278
16.2521 | -104.0472
16.2558 | -104.0714
16.2596 | -104.093
16.2633 | -104.1108
16.2671 | -104.1249
16.2708 | -104.1353
16.2746 | -104.141
16.2783 | -104.146
16.2821 | -104.146
16.2858 | -104.1408
16.2896 | -104.137
16.2933 | -104.1322
16.2971 | -104.1248
16.3008 | -104.1138
16.3046 | -104.1005
16.3083 | -104.0884
16.3121 | -104.0717
16.3158 | -104.0518
16.3196 | -104.0338
16.3233 | -104.0163
16.3271 | -103.9985
16.3308 | -103.979
16.3346 | -103.9569
16.3383 | -103.9356
16.3421 | -103.9164
16.3458 | -103.8995
16.3496 | -103.884
16.3533 | -103.8707
16.3571 | -103.8605
16.3608 | -103.8507
16.3646 | -103.8427
16.3683 | -103.8369
16.3721 | -103.8346
16.3758 | -103.8367
16.3796 | -103.8419
16.3833 | -103.8506
16.3871 | -103.8608
16.3908 | -103.8754
16.3946 | -103.8938
16.3983 | -103.9166
16.4021 | -103.9408
16.4058 | -103.9681
16.4096 | -103.9967
16.4133 | -104.0294
16.4171 | -104.0644
16.4208 | -104.1025
16.4246 | -104.143
16.4283 | -104.1832
16.4321 | -104.2222
16.4358 | -104.2605
16.4396 | -104.2988
16.4433 | -104.3387
16.4471 | -104.3773
16.4508 | -104.4112
16.4546 | -104.4424
16.4583 | -104.4689
16.4621 | -104.4931
16.4658 | -104.5175
16.4696 | -104.5387
16.4733 | -104.5568
16.4771 | -104.5747
16.4808 | -104.5854
16.4846 | -104.5909
16.4883 | -104.593
16.4921 | -104.5886
16.4959 | -104.5814
16.4996 | -104.5731
16.5034 | -104.5629
16.5071 | -104.5466
16.5109 | -104.528
16.5146 | -104.5025
16.5184 | -104.4761
16.5221 | -104.4549
16.5259 | -104.4321
16.5296 | -104.4052
16.5334 | -104.3763
16.5371 | -104.3475
16.5409 | -104.3173
16.5446 | -104.2888
16.5484 | -104.2604
16.5521 | -104.2309
16.5559 | -104.2011
16.5596 | -104.1735
16.5634 | -104.1483
16.5671 | -104.1242
16.5709 | -104.103
16.5746 | -104.0828
16.5784 | -104.0656
16.5821 | -104.0495
16.5859 | -104.0363
16.5896 | -104.026
16.5934 | -104.016
16.5971 | -104.0107
16.6009 | -104.011
16.6046 | -104.0138
16.6084 | -104.0177
16.6121 | -104.0249
16.6159 | -104.0342
16.6196 | -104.0423
16.6234 | -104.0542
16.6271 | -104.0722
16.6309 | -104.0922
16.6346 | -104.1116
16.6384 | -104.1319
16.6421 | -104.1527
16.6459 | -104.1749
16.6496 | -104.1996
16.6534 | -104.2238
16.6571 | -104.2468
16.6609 | -104.2688
16.6646 | -104.2905
16.6684 | -104.3098
16.6721 | -104.3253
16.6759 | -104.337
16.6796 | -104.3443
16.6834 | -104.3502
16.6871 | -104.3545
16.6909 | -104.3554
16.6946 | -104.3507
16.6984 | -104.3437
16.7021 | -104.331
16.7059 | -104.3151
16.7096 | -104.2997
16.7134 | -104.2805
16.7171 | -104.2618
16.7209 | -104.2419
16.7246 | -104.2201
16.7284 | -104.1989
16.7321 | -104.1766
16.7359 | -104.1527
16.7397 | -104.1319
16.7434 | -104.1168
16.7472 | -104.1014
16.7509 | -104.0905
16.7547 | -104.0831
16.7584 | -104.0761
16.7622 | -104.0698
16.7659 | -104.0682
16.7697 | -104.0696
16.7734 | -104.0963
16.7772 | -104.1066
16.7809 | -104.1211
16.7847 | -104.1365
16.7884 | -104.1557
16.7922 | -104.181
16.7959 | -104.2086
16.7997 | -104.2363
16.8034 | -104.2672
16.8072 | -104.3011
16.8109 | -104.3353
16.8147 | -104.3724
16.8184 | -104.411
16.8222 | -104.4514
16.8259 | -104.4924
16.8297 | -104.534
16.8334 | -104.5752
16.8372 | -104.6146
16.8409 | -104.6547
16.8447 | -104.6939
16.8484 | -104.7328
16.8522 | -104.7701
16.8559 | -104.8073
16.8597 | -104.8474
16.8634 | -104.8851
16.8672 | -104.9172
16.8709 | -104.9472
16.8747 | -104.9788
16.8784 | -105.0105
16.8822 | -105.0395
16.8859 | -105.0662
16.8897 | -105.0887
16.8934 | -105.1066
16.8972 | -105.1233
16.9009 | -105.1398
16.9047 | -105.1515
16.9084 | -105.1646
16.9122 | -105.18
16.9159 | -105.1945
16.9197 | -105.207
16.9234 | -105.2149
16.9272 | -105.22
16.9309 | -105.2218
16.9347 | -105.2227
16.9384 | -105.2191
16.9422 | -105.2163
16.9459 | -105.2112
16.9497 | -105.2022
16.9534 | -105.1928
16.9572 | -105.1801
16.9609 | -105.1642
16.9647 | -105.1463
16.9684 | -105.1273
16.9722 | -105.1078
16.9759 | -105.0847
16.9797 | -105.057
16.9835 | -105.0255
16.9872 | -104.9934
16.991 | -104.9596
16.9947 | -104.9234
16.9985 | -104.8866
17.0022 | -104.8506
17.006 | -104.8115
17.0097 | -104.7698
17.0135 | -104.7284
17.0172 | -104.6862
17.021 | -104.6436
17.0247 | -104.6016
17.0285 | -104.5614
17.0322 | -104.5206
17.036 | -104.4794
17.0397 | -104.443
17.0435 | -104.4082
17.0472 | -104.3747
17.051 | -104.343
17.0547 | -104.3134
17.0585 | -104.2859
17.0622 | -104.2604
17.066 | -104.2352
17.0697 | -104.2112
17.0735 | -104.193
17.0772 | -104.1771
17.081 | -104.1637
17.0847 | -104.152
17.0885 | -104.1446
17.0922 | -104.1384
17.096 | -104.1347
17.0997 | -104.1319
17.1035 | -104.1321
17.1072 | -104.1308
17.111 | -104.1302
17.1147 | -104.1331
17.1185 | -104.1363
17.1222 | -104.1404
17.126 | -104.1459
17.1297 | -104.1524
17.1335 | -104.1598
17.1372 | -104.1692
17.141 | -104.1792
17.1447 | -104.1873
17.1485 | -104.1971
17.1522 | -104.2049
17.156 | -104.2116
17.1597 | -104.2169
17.1635 | -104.2214
17.1672 | -104.2271
17.171 | -104.2336
17.1747 | -104.2408
17.1785 | -104.2458
17.1822 | -104.2508
17.186 | -104.2545
17.1897 | -104.2602
17.1935 | -104.2642
17.1972 | -104.2676
17.201 | -104.2731
17.2047 | -104.2832
17.2085 | -104.2939
17.2122 | -104.302
17.216 | -104.3125
17.2197 | -104.3249
17.2235 | -104.3385
17.2272 | -104.3551
17.231 | -104.372
17.2348 | -104.3885
17.2385 | -104.4054
17.2423 | -104.4213
17.246 | -104.4401
17.2498 | -104.4594
17.2535 | -104.481
17.2573 | -104.5027
17.261 | -104.5229
17.2648 | -104.542
17.2685 | -104.5574
17.2723 | -104.5689
17.276 | -104.5816
17.2798 | -104.5972
17.2835 | -104.6085
17.2873 | -104.6205
17.291 | -104.6316
17.2948 | -104.6406
17.2985 | -104.6476
17.3023 | -104.6515
17.306 | -104.6548
17.3098 | -104.6611
17.3135 | -104.6594
17.3173 | -104.6662
17.321 | -104.6661
17.3248 | -104.666
17.3285 | -104.6643
17.3323 | -104.6626
17.336 | -104.6593
17.3398 | -104.6559
17.3435 | -104.6498
17.3473 | -104.6455
17.351 | -104.6414
17.3548 | -104.6374
17.3585 | -104.6331
17.3623 | -104.6269
17.366 | -104.6181
17.3698 | -104.6083
17.3735 | -104.5985
17.3773 | -104.5884
17.381 | -104.5772
17.3848 | -104.5669
17.3885 | -104.5563
17.3923 | -104.5429
17.396 | -104.5293
17.3998 | -104.5167
17.4035 | -104.5027
17.4073 | -104.4901
17.411 | -104.4779
17.4148 | -104.4627
17.4185 | -104.447
17.4223 | -104.4317
17.426 | -104.4161
17.4298 | -104.3982
17.4335 | -104.3806
17.4373 | -104.3621
17.441 | -104.3449
17.4448 | -104.3302
17.4485 | -104.3145
17.4523 | -104.2949
17.456 | -104.2764
17.4598 | -104.2616
17.4635 | -104.2472
17.4673 | -104.2333
17.471 | -104.221
17.4748 | -104.2073
17.4785 | -104.1955
17.4823 | -104.1854
17.4861 | -104.1771
17.4898 | -104.1691
17.4936 | -104.1612
17.4973 | -104.1573
17.5011 | -104.1553
17.5048 | -104.1538
17.5086 | -104.153
17.5123 | -104.1534
17.5161 | -104.1586
17.5198 | -104.1679
17.5236 | -104.1761
17.5273 | -104.1815
17.5311 | -104.1902
17.5348 | -104.2008
17.5386 | -104.2129
17.5423 | -104.2275
17.5461 | -104.2434
17.5498 | -104.2612
17.5536 | -104.2794
17.5573 | -104.2971
17.5611 | -104.3143
17.5648 | -104.3302
17.5686 | -104.3476
17.5723 | -104.3689
17.5761 | -104.3881
17.5798 | -104.4083
17.5836 | -104.4239
17.5873 | -104.4385
17.5911 | -104.4545
17.5948 | -104.4692
17.5986 | -104.482
17.6023 | -104.4922
17.6061 | -104.5035
17.6098 | -104.5108
17.6136 | -104.5161
17.6173 | -104.5221
17.6211 | -104.5259
17.6248 | -104.526
17.6286 | -104.5241
17.6323 | -104.5209
17.6361 | -104.5163
17.6398 | -104.512
17.6436 | -104.5095
17.6473 | -104.5046
17.6511 | -104.4986
17.6548 | -104.4911
17.6586 | -104.4818
17.6623 | -104.4735
17.6661 | -104.4695
17.6698 | -104.4674
17.6736 | -104.464
17.6773 | -104.4613
17.6811 | -104.4603
17.6848 | -104.4578
17.6886 | -104.457
17.6923 | -104.458
17.6961 | -104.4578
17.6998 | -104.458
17.7036 | -104.4586
17.7073 | -104.4597
17.7111 | -104.4601
17.7148 | -104.4627
17.7186 | -104.4689
17.7223 | -104.4745
17.7261 | -104.482
17.7299 | -104.4904
17.7336 | -104.4996
17.7374 | -104.5085
17.7411 | -104.5176
17.7449 | -104.5271
17.7486 | -104.54
17.7524 | -104.5527
17.7561 | -104.5657
17.7599 | -104.579
17.7636 | -104.5931
17.7674 | -104.6085
17.7711 | -104.623
17.7749 | -104.6379
17.7786 | -104.6542
17.7824 | -104.6698
17.7861 | -104.6883
17.7899 | -104.7071
17.7936 | -104.7241
17.7974 | -104.7454
17.8011 | -104.7667
17.8049 | -104.7887
17.8086 | -104.8143
17.8124 | -104.8417
17.8161 | -104.8685
17.8199 | -104.8934
17.8236 | -104.9176
17.8274 | -104.9443
17.8311 | -104.9717
17.8349 | -104.9997
17.8386 | -105.0257
17.8424 | -105.05
17.8461 | -105.0728
17.8499 | -105.0945
17.8536 | -105.1115
17.8574 | -105.1315
17.8611 | -105.1488
17.8649 | -105.1662
17.8686 | -105.1825
17.8724 | -105.1972
17.8761 | -105.2114
17.8799 | -105.2241
17.8836 | -105.231
17.8874 | -105.2356
17.8911 | -105.2397
17.8949 | -105.24
17.8986 | -105.2372
17.9024 | -105.2336
17.9061 | -105.2279
17.9099 | -105.2243
17.9136 | -105.2212
17.9174 | -105.2154
17.9211 | -105.2026
17.9249 | -105.189
17.9286 | -105.1783
17.9324 | -105.1668
17.9361 | -105.1555
17.9399 | -105.1442
17.9436 | -105.131
17.9474 | -105.1135
17.9511 | -105.0976
17.9549 | -105.0842
17.9586 | -105.0748
17.9624 | -105.0661
17.9661 | -105.0581
17.9699 | -105.0498
17.9737 | -105.04
17.9774 | -105.0336
17.9812 | -105.0287
17.9849 | -105.0223
17.9887 | -105.0168
17.9924 | -105.0156
17.9962 | -105.0164
17.9999 | -105.0169
18.0037 | -105.0161
18.0074 | -105.017
18.0112 | -105.0206
18.0149 | -105.0271
18.0187 | -105.034
18.0224 | -105.039
18.0262 | -105.0447
18.0299 | -105.0537
18.0337 | -105.0626
18.0374 | -105.0728
18.0412 | -105.081
18.0449 | -105.09
18.0487 | -105.1014
18.0524 | -105.1097
18.0562 | -105.1222
18.0599 | -105.1333
18.0637 | -105.1416
18.0674 | -105.1517
18.0712 | -105.1641
18.0749 | -105.1755
18.0787 | -105.1879
18.0824 | -105.2004
18.0862 | -105.2099
18.0899 | -105.2178
18.0937 | -105.2264
18.0974 | -105.2333
18.1012 | -105.2416
18.1049 | -105.2465
18.1087 | -105.2489
18.1124 | -105.2502
18.1162 | -105.2495
18.1199 | -105.2481
18.1237 | -105.2482
18.1274 | -105.2466
18.1312 | -105.243
18.1349 | -105.2363
18.1387 | -105.2294
18.1424 | -105.224
18.1462 | -105.2192
18.1499 | -105.2102
18.1537 | -105.2026
18.1574 | -105.1961
18.1612 | -105.1867
18.1649 | -105.177
18.1687 | -105.1651
18.1724 | -105.1523
18.1762 | -105.1407
18.1799 | -105.1325
18.1837 | -105.1215
18.1874 | -105.1067
18.1912 | -105.0925
18.1949 | -105.0774
18.1987 | -105.063
18.2024 | -105.0494
18.2062 | -105.0384
18.2099 | -105.0253
18.2137 | -105.011
18.2174 | -104.9946
18.2212 | -104.9787
18.225 | -104.9632
18.2287 | -104.9478
18.2325 | -104.934
18.2362 | -104.9206
18.24 | -104.9068
18.2437 | -104.8955
18.2475 | -104.8866
18.2512 | -104.8773
18.255 | -104.869
18.2587 | -104.861
18.2625 | -104.8557
18.2662 | -104.8525
18.27 | -104.8499
18.2737 | -104.8471
18.2775 | -104.8454
18.2812 | -104.8479
18.285 | -104.8524
18.2887 | -104.8561
18.2925 | -104.8586
18.2962 | -104.862
18.3 | -104.8649
18.3037 | -104.869
18.3075 | -104.874
18.3112 | -104.8787
18.315 | -104.8779
18.3187 | -104.8807
18.3225 | -104.8863
18.3262 | -104.8928
18.33 | -104.8952
18.3337 | -104.8932
18.3375 | -104.8865
18.3412 | -104.8838
18.345 | -104.8826
18.3487 | -104.8803
18.3525 | -104.876
18.3562 | -104.868
18.36 | -104.8568
18.3637 | -104.8474
18.3675 | -104.836
18.3712 | -104.8249
18.375 | -104.8142
18.3787 | -104.8004
18.3825 | -104.7816
18.3862 | -104.7628
18.39 | -104.7443
18.3937 | -104.7271
18.3975 | -104.7011
18.4012 | -104.681
18.405 | -104.6614
18.4087 | -104.639
18.4125 | -104.6171
18.4162 | -104.595
18.42 | -104.5744
18.4237 | -104.555
18.4275 | -104.5335
18.4312 | -104.5109
18.435 | -104.4915
18.4387 | -104.4742
18.4425 | -104.4573
18.4462 | -104.4405
18.45 | -104.4267
18.4537 | -104.4151
18.4575 | -104.4029
18.4612 | -104.3919
18.465 | -104.3825
18.4688 | -104.3732
18.4725 | -104.3665
18.4763 | -104.3604
18.48 | -104.3527
18.4838 | -104.3459
18.4875 | -104.3429
18.4913 | -104.3419
18.495 | -104.3401
18.4988 | -104.3363
18.5025 | -104.3304
18.5063 | -104.3208
18.51 | -104.3117
18.5138 | -104.3046
18.5175 | -104.2959
18.5213 | -104.2859
18.525 | -104.2722
18.5288 | -104.2584
18.5325 | -104.2437
18.5363 | -104.2304
18.54 | -104.2173
18.5438 | -104.1999
18.5475 | -104.1796
18.5513 | -104.1583
18.555 | -104.1378
18.5588 | -104.1159
18.5625 | -104.0937
18.5663 | -104.0723
18.57 | -104.0477
18.5738 | -104.022
18.5775 | -103.9972
18.5813 | -103.9695
18.585 | -103.9407
18.5888 | -103.9122
18.5925 | -103.8875
18.5963 | -103.8657
18.6 | -103.8456
18.6038 | -103.823
18.6075 | -103.8
18.6113 | -103.7784
18.615 | -103.76
18.6188 | -103.745
18.6225 | -103.7281
18.6263 | -103.7096
18.63 | -103.6963
18.6338 | -103.6849
18.6375 | -103.6768
18.6413 | -103.674
18.645 | -103.6725
18.6488 | -103.6714
18.6525 | -103.6712
18.6563 | -103.6744
18.66 | -103.6828
18.6638 | -103.6944
18.6675 | -103.7065
18.6713 | -103.7179
18.675 | -103.7308
18.6788 | -103.7471
18.6825 | -103.7664
18.6863 | -103.785
18.69 | -103.8049
18.6938 | -103.8277
18.6975 | -103.8532
18.7013 | -103.8753
18.705 | -103.8969
18.7088 | -103.9187
18.7126 | -103.9419
18.7163 | -103.9643
18.7201 | -103.9837
18.7238 | -104.0034
18.7276 | -104.0222
18.7313 | -104.0403
18.7351 | -104.0568
18.7388 | -104.0696
18.7426 | -104.0829
18.7463 | -104.097
18.7501 | -104.1087
18.7538 | -104.1181
18.7576 | -104.1243
18.7613 | -104.1306
18.7651 | -104.1363
18.7688 | -104.1419
18.7726 | -104.1485
18.7763 | -104.1508
18.7801 | -104.1453
18.7838 | -104.1383
18.7876 | -104.1318
18.7913 | -104.1267
18.7951 | -104.1172
18.7988 | -104.1043
18.8026 | -104.0898
18.8063 | -104.0736
18.8101 | -104.0573
18.8138 | -104.0399
18.8176 | -104.02
18.8213 | -103.999
18.8251 | -103.9782
18.8288 | -103.9587
18.8326 | -103.9377
18.8363 | -103.9164
18.8401 | -103.8937
18.8438 | -103.8701
18.8476 | -103.8443
18.8513 | -103.8194
18.8551 | -103.7941
18.8588 | -103.7679
18.8626 | -103.743
18.8663 | -103.7143
18.8701 | -103.6832
18.8738 | -103.6585
18.8776 | -103.6345
18.8813 | -103.6097
18.8851 | -103.5844
18.8888 | -103.5597
18.8926 | -103.532
18.8963 | -103.4985
18.9001 | -103.4658
18.9038 | -103.4302
18.9076 | -103.3935
18.9113 | -103.3565
18.9151 | -103.315
18.9188 | -103.2719
18.9226 | -103.2292
18.9263 | -103.1845
18.9301 | -103.1395
18.9338 | -103.0958
18.9376 | -103.0419
18.9413 | -102.9974
18.9451 | -102.9548
18.9488 | -102.9105
18.9526 | -102.8652
18.9564 | -102.8164
18.9601 | -102.7655
18.9639 | -102.7167
18.9676 | -102.6699
18.9714 | -102.6271
18.9751 | -102.5841
18.9789 | -102.5429
18.9826 | -102.5035
18.9864 | -102.4658
18.9901 | -102.4317
18.9939 | -102.3992
18.9976 | -102.3687
19.0014 | -102.3412
19.0051 | -102.3165
19.0089 | -102.2946
19.0126 | -102.2742
19.0164 | -102.2558
19.0201 | -102.2393
19.0239 | -102.2239
19.0276 | -102.2124
19.0314 | -102.2017
19.0351 | -102.1939
19.0389 | -102.1906
19.0426 | -102.1892
19.0464 | -102.1913
19.0501 | -102.1972
19.0539 | -102.2021
19.0576 | -102.2072
19.0614 | -102.213
19.0651 | -102.2202
19.0689 | -102.2294
19.0726 | -102.2398
19.0764 | -102.2495
19.0801 | -102.2573
19.0839 | -102.2586
19.0876 | -102.2595
19.0914 | -102.2656
19.0951 | -102.2737
19.0989 | -102.2778
19.1026 | -102.2786
19.1064 | -102.2775
19.1101 | -102.2737
19.1139 | -102.2715
19.1176 | -102.27
19.1214 | -102.2639
19.1251 | -102.2526
19.1289 | -102.2366
19.1326 | -102.2161
19.1364 | -102.1981
19.1401 | -102.1793
19.1439 | -102.1541
19.1476 | -102.1259
19.1514 | -102.0947
19.1551 | -102.0618
19.1589 | -102.0293
19.1626 | -101.997
19.1664 | -101.9643
19.1701 | -101.9304
19.1739 | -101.8948
19.1776 | -101.8588
19.1814 | -101.8231
19.1851 | -101.7893
19.1889 | -101.7576
19.1926 | -101.726
19.1964 | -101.6943
19.2002 | -101.6645
19.2039 | -101.6374
19.2077 | -101.6121
19.2114 | -101.5895
19.2152 | -101.5706
19.2189 | -101.5526
19.2227 | -101.5367
19.2264 | -101.5229
19.2302 | -101.5166
19.2339 | -101.5153
19.2377 | -101.5141
19.2414 | -101.5146
19.2452 | -101.5166
19.2489 | -101.5177
19.2527 | -101.5185
19.2564 | -101.5224
19.2602 | -101.5288
19.2639 | -101.5378
19.2677 | -101.5477
19.2714 | -101.5554
19.2752 | -101.5628
19.2789 | -101.5713
19.2827 | -101.5823
19.2864 | -101.5921
19.2902 | -101.6021
19.2939 | -101.6123
19.2977 | -101.6175
19.3014 | -101.6167
19.3052 | -101.6179
19.3089 | -101.616
19.3127 | -101.6126
19.3164 | -101.6091
19.3202 | -101.5975
19.3239 | -101.5783
19.3277 | -101.5604
19.3314 | -101.5411
19.3352 | -101.5219
19.3389 | -101.5034
19.3427 | -101.4799
19.3464 | -101.4494
19.3502 | -101.4173
19.3539 | -101.3878
19.3577 | -101.3629
19.3614 | -101.3366
19.3652 | -101.3031
19.3689 | -101.2658
19.3727 | -101.2296
19.3764 | -101.1947
19.3802 | -101.1607
19.3839 | -101.1265
19.3877 | -101.0883
19.3914 | -101.0479
19.3952 | -101.0092
19.3989 | -100.9745
19.4027 | -100.9417
19.4064 | -100.9103
19.4102 | -100.8793
19.4139 | -100.8485
19.4177 | -100.8197
19.4214 | -100.7939
19.4252 | -100.7698
19.4289 | -100.749
19.4327 | -100.7318
19.4364 | -100.7173
19.4402 | -100.7048
19.444 | -100.6941
19.4477 | -100.6869
19.4515 | -100.6831
19.4552 | -100.684
19.459 | -100.6902
19.4627 | -100.697
19.4665 | -100.7059
19.4702 | -100.7204
19.474 | -100.7395
19.4777 | -100.7871
19.4815 | -100.8126
19.4852 | -100.8425
19.489 | -100.877
19.4927 | -100.9146
19.4965 | -100.955
19.5002 | -100.9983
19.504 | -101.0458
19.5077 | -101.0936
19.5115 | -101.1434
19.5152 | -101.1947
19.519 | -101.249
19.5227 | -101.3055
19.5265 | -101.3649
19.5302 | -101.426
19.534 | -101.4844
19.5377 | -101.5414
19.5415 | -101.5958
19.5452 | -101.6511
19.549 | -101.7089
19.5527 | -101.7661
19.5565 | -101.8197
19.5602 | -101.8645
19.564 | -101.9024
19.5677 | -101.9358
19.5715 | -101.9687
19.5752 | -102.0075
19.579 | -102.0459
19.5827 | -102.0716
19.5865 | -102.0871
19.5902 | -102.0939
19.594 | -102.0967
19.5977 | -102.1059
19.6015 | -102.1166
19.6052 | -102.1177
19.609 | -102.1124
19.6127 | -102.0994
19.6165 | -102.0792
19.6202 | -102.0585
19.624 | -102.0399
19.6277 | -102.0176
19.6315 | -101.9875
19.6352 | -101.9496
19.639 | -101.9035
19.6427 | -101.856
19.6465 | -101.8134
19.6502 | -101.7729
19.654 | -101.7272
19.6577 | -101.673
19.6615 | -101.6131
19.6652 | -101.5544
19.669 | -101.5037
19.6727 | -101.4583
19.6765 | -101.4118
19.6802 | -101.3628
19.684 | -101.3138
19.6877 | -101.2646
19.6915 | -101.2194
19.6953 | -101.1774
19.699 | -101.137
19.7028 | -101.0985
19.7065 | -101.0615
19.7103 | -101.0247
19.714 | -100.9897
19.7178 | -100.9579
19.7215 | -100.9295
19.7253 | -100.9045
19.729 | -100.8795
19.7328 | -100.855
19.7365 | -100.8318
19.7403 | -100.8126
19.744 | -100.7972
19.7478 | -100.7839
19.7515 | -100.772
19.7553 | -100.7621
19.759 | -100.7545
19.7628 | -100.7498
19.7665 | -100.7473
19.7703 | -100.7478
19.774 | -100.7517
19.7778 | -100.7594
19.7815 | -100.7705
19.7853 | -100.7823
19.789 | -100.7984
19.7928 | -100.818
19.7965 | -100.8416
19.8003 | -100.8666
19.804 | -100.8944
19.8078 | -100.9276
19.8115 | -100.9636
19.8153 | -101.0017
19.819 | -101.0434
19.8228 | -101.0893
19.8265 | -101.1363
19.8303 | -101.1844
19.834 | -101.2342
19.8378 | -101.2884
19.8415 | -101.3488
19.8453 | -101.414
19.849 | -101.4792
19.8528 | -101.5415
19.8565 | -101.6
19.8603 | -101.6576
19.864 | -101.7209
19.8678 | -101.7878
19.8715 | -101.8534
19.8753 | -101.9151
19.879 | -101.969
19.8828 | -102.019
19.8865 | -102.0703
19.8903 | -102.1203
19.894 | -102.1691
19.8978 | -102.215
19.9015 | -102.2568
19.9053 | -102.2951
19.909 | -102.3315
19.9128 | -102.3679
19.9165 | -102.4059
19.9203 | -102.4458
19.924 | -102.4798
19.9278 | -102.5023
19.9315 | -102.5196
19.9353 | -102.5384
19.9391 | -102.5635
19.9428 | -102.5942
19.9466 | -102.6165
19.9503 | -102.6201
19.9541 | -102.6098
19.9578 | -102.6013
19.9616 | -102.6006
19.9653 | -102.5991
19.9691 | -102.5926
19.9728 | -102.5796
19.9766 | -102.5622
19.9803 | -102.5398
19.9841 | -102.5171
19.9878 | -102.4908
19.9916 | -102.4635
19.9953 | -102.4362
19.9991 | -102.4047
20.0028 | -102.3675
20.0066 | -102.3265
20.0103 | -102.2856
20.0141 | -102.2427
20.0178 | -102.175
20.0216 | -102.1277
20.0253 | -102.0778
20.0291 | -102.026
20.0328 | -101.9722
20.0366 | -101.9176
20.0403 | -101.8609
20.0441 | -101.8032
20.0478 | -101.7458
20.0516 | -101.6894
20.0553 | -101.6327
20.0591 | -101.5748
20.0628 | -101.5171
20.0666 | -101.4583
20.0703 | -101.4017
20.0741 | -101.3486
20.0778 | -101.2969
20.0816 | -101.2472
20.0853 | -101.2001
20.0891 | -101.1562
20.0928 | -101.1147
20.0966 | -101.0761
20.1003 | -101.0425
20.1041 | -101.0117
20.1078 | -100.9832
20.1116 | -100.9592
20.1153 | -100.94
20.1191 | -100.9244
20.1228 | -100.9112
20.1266 | -100.9024
20.1303 | -100.8978
20.1341 | -100.8974
20.1378 | -100.9021
20.1416 | -100.9103
20.1453 | -100.9235
20.1491 | -100.9401
20.1528 | -100.9606
20.1566 | -100.9856
20.1603 | -101.0134
20.1641 | -101.0442
20.1678 | -101.0782
20.1716 | -101.1142
20.1753 | -101.1527
20.1791 | -101.193
20.1828 | -101.2363
20.1866 | -101.2812
20.1904 | -101.3268
20.1941 | -101.3728
20.1979 | -101.4184
20.2016 | -101.4624
20.2054 | -101.5045
20.2091 | -101.5475
20.2129 | -101.5906
20.2166 | -101.6323
20.2204 | -101.673
20.2241 | -101.7103
20.2279 | -101.7411
20.2316 | -101.7693
20.2354 | -101.7971
20.2391 | -101.8239
20.2429 | -101.8434
20.2466 | -101.8545
20.2504 | -101.861
20.2541 | -101.8644
20.2579 | -101.8636
20.2616 | -101.857
20.2654 | -101.8474
20.2691 | -101.8318
20.2729 | -101.8113
20.2766 | -101.7857
20.2804 | -101.7562
20.2841 | -101.7231
20.2879 | -101.6887
20.2916 | -101.6523
20.2954 | -101.6123
20.2991 | -101.5694
20.3029 | -101.526
20.3066 | -101.4805
20.3104 | -101.435
20.3141 | -101.3899
20.3179 | -101.3442
20.3216 | -101.2985
20.3254 | -101.2544
20.3291 | -101.2152
20.3329 | -101.1779
20.3366 | -101.1423
20.3404 | -101.1085
20.3441 | -101.0765
20.3479 | -101.0492
20.3516 | -101.0267
20.3554 | -101.0096
20.3591 | -100.9969
20.3629 | -100.9859
20.3666 | -100.9785
20.3704 | -100.9753
20.3741 | -100.9769
20.3779 | -100.9824
20.3816 | -100.992
20.3854 | -101.0073
20.3891 | -101.0263
20.3929 | -101.0497
20.3966 | -101.0784
20.4004 | -101.1112
20.4041 | -101.1468
20.4079 | -101.1859
20.4116 | -101.2303
20.4154 | -101.2757
20.4191 | -101.3237
20.4229 | -101.3757
20.4266 | -101.4289
20.4304 | -101.4854
20.4342 | -101.5425
20.4379 | -101.5992
20.4417 | -101.6577
20.4454 | -101.7156
20.4492 | -101.7732
20.4529 | -101.8292
20.4567 | -101.8839
20.4604 | -101.939
20.4642 | -101.9915
20.4679 | -102.0383
20.4717 | -102.0793
20.4754 | -102.1169
20.4792 | -102.1521
20.4829 | -102.1846
20.4867 | -102.2126
20.4904 | -102.2356
20.4942 | -102.2544
20.4979 | -102.2651
20.5017 | -102.2691
20.5054 | -102.2662
20.5092 | -102.2587
20.5129 | -102.246
20.5167 | -102.2276
20.5204 | -102.206
20.5242 | -102.1805
20.5279 | -102.1494
20.5317 | -102.1139
20.5354 | -102.0773
20.5392 | -102.0382
20.5429 | -101.9973
20.5467 | -101.9548
20.5504 | -101.9127
20.5542 | -101.8706
20.5579 | -101.8276
20.5617 | -101.7912
20.5654 | -101.7532
20.5692 | -101.7158
20.5729 | -101.6777
20.5767 | -101.6414
20.5804 | -101.6064
20.5842 | -101.5748
20.5879 | -101.5467
20.5917 | -101.5231
20.5954 | -101.5026
20.5992 | -101.4852
20.6029 | -101.47
20.6067 | -101.4578
20.6104 | -101.4474
20.6142 | -101.44
20.6179 | -101.4358
20.6217 | -101.434
20.6254 | -101.4358
20.6292 | -101.4416
20.6329 | -101.4494
20.6367 | -101.4589
20.6404 | -101.4717
20.6442 | -101.4869
20.6479 | -101.5021
20.6517 | -101.5155
20.6554 | -101.5296
20.6592 | -101.5468
20.6629 | -101.5653
20.6667 | -101.5837
20.6704 | -101.6028
20.6742 | -101.6223
20.6779 | -101.6404
20.6817 | -101.658
20.6855 | -101.6741
20.6892 | -101.6878
20.693 | -101.6967
20.6967 | -101.7009
20.7005 | -101.7046
20.7042 | -101.7086
20.708 | -101.7126
20.7117 | -101.7145
20.7155 | -101.7122
20.7192 | -101.7061
20.723 | -101.6971
20.7267 | -101.6865
20.7305 | -101.6741
20.7342 | -101.6585
20.738 | -101.6406
20.7417 | -101.6236
20.7455 | -101.6071
20.7492 | -101.5877
20.753 | -101.5643
20.7567 | -101.5415
20.7605 | -101.519
20.7642 | -101.4951
20.768 | -101.4685
20.7717 | -101.4381
20.7755 | -101.4075
20.7792 | -101.3777
20.783 | -101.35
20.7867 | -101.3246
20.7905 | -101.2991
20.7942 | -101.2736
20.798 | -101.2501
20.8017 | -101.2284
20.8055 | -101.2086
20.8092 | -101.1913
20.813 | -101.1761
20.8167 | -101.1635
20.8205 | -101.1531
20.8242 | -101.1452
20.828 | -101.1377
20.8317 | -101.1322
20.8355 | -101.1313
20.8392 | -101.1335
20.843 | -101.1382
20.8467 | -101.1459
20.8505 | -101.1571
20.8542 | -101.1704
20.858 | -101.1866
20.8617 | -101.2062
20.8655 | -101.2294
20.8692 | -101.2547
20.873 | -101.2822
20.8767 | -101.3136
20.8805 | -101.3446
20.8842 | -101.3788
20.888 | -101.4142
20.8917 | -101.452
20.8955 | -101.4913
20.8992 | -101.5333
20.903 | -101.5766
20.9067 | -101.6231
20.9105 | -101.6697
20.9142 | -101.7148
20.918 | -101.7614
20.9217 | -101.8089
20.9255 | -101.855
20.9293 | -101.904
20.933 | -101.9542
20.9368 | -102.0017
20.9405 | -102.0488
20.9443 | -102.0974
20.948 | -102.1488
20.9518 | -102.1981
20.9555 | -102.2421
20.9593 | -102.2852
20.963 | -102.3265
20.9668 | -102.3626
20.9705 | -102.3962
20.9743 | -102.4297
20.978 | -102.4615
20.9818 | -102.4895
20.9855 | -102.5156
20.9893 | -102.5399
20.993 | -102.5665
20.9968 | -102.589
21.0005 | -102.6039
21.0043 | -102.615
21.008 | -102.6224
21.0118 | -102.626
21.0155 | -102.6273
21.0193 | -102.6268
21.023 | -102.6236
21.0268 | -102.6192
21.0305 | -102.6135
21.0343 | -102.6038
21.038 | -102.5877
21.0418 | -102.5716
21.0455 | -102.5568
21.0493 | -102.5407
21.053 | -102.5218
21.0568 | -102.5022
21.0605 | -102.4843
21.0643 | -102.47
21.068 | -102.4556
21.0718 | -102.4385
21.0755 | -102.4167
21.0793 | -102.393
21.083 | -102.3704
21.0868 | -102.3521
21.0905 | -102.3349
21.0943 | -102.3174
21.098 | -102.2977
21.1018 | -102.2844
21.1055 | -102.2675
21.1093 | -102.252
21.113 | -102.2379
21.1168 | -102.2245
21.1205 | -102.2133
21.1243 | -102.2044
21.128 | -102.1986
21.1318 | -102.1947
21.1355 | -102.19
21.1393 | -102.1866
21.143 | -102.1844
21.1468 | -102.184
21.1505 | -102.1836
21.1543 | -102.1844
21.158 | -102.1883
21.1618 | -102.1927
21.1655 | -102.1976
21.1693 | -102.2043
21.1731 | -102.2132
21.1768 | -102.2223
21.1806 | -102.2334
21.1843 | -102.2439
21.1881 | -102.2562
21.1918 | -102.269
21.1956 | -102.2839
21.1993 | -102.2986
21.2031 | -102.3135
21.2068 | -102.3271
21.2106 | -102.3431
21.2143 | -102.3619
21.2181 | -102.3795
21.2218 | -102.3939
21.2256 | -102.4037
21.2293 | -102.4126
21.2331 | -102.4225
21.2368 | -102.4324
21.2406 | -102.4407
21.2443 | -102.4496
21.2481 | -102.4534
21.2518 | -102.4523
21.2556 | -102.4514
21.2593 | -102.449
21.2631 | -102.4459
21.2668 | -102.4467
21.2706 | -102.4446
21.2743 | -102.4338
21.2781 | -102.4179
21.2818 | -102.4011
21.2856 | -102.3843
21.2893 | -102.3689
21.2931 | -102.3504
21.2968 | -102.3297
21.3006 | -102.3074
21.3043 | -102.2814
21.3081 | -102.254
21.3118 | -102.2284
21.3156 | -102.2025
21.3193 | -102.1757
21.3231 | -102.1456
21.3268 | -102.1082
21.3306 | -102.0707
21.3343 | -102.038
21.3381 | -102.0092
21.3418 | -101.9785
21.3456 | -101.9465
21.3493 | -101.915
21.3531 | -101.8832
21.3568 | -101.8536
21.3606 | -101.8264
21.3643 | -101.7984
21.3681 | -101.7718
21.3718 | -101.7462
21.3756 | -101.726
21.3793 | -101.706
21.3831 | -101.685
21.3868 | -101.6645
21.3906 | -101.6466
21.3943 | -101.6301
21.3981 | -101.6146
21.4018 | -101.6009
21.4056 | -101.589
21.4093 | -101.5787
21.4131 | -101.5701
21.4169 | -101.5627
21.4206 | -101.556
21.4244 | -101.5515
21.4281 | -101.5477
21.4319 | -101.5458
21.4356 | -101.5474
21.4394 | -101.5519
21.4431 | -101.5588
21.4469 | -101.5668
21.4506 | -101.5759
21.4544 | -101.5892
21.4581 | -101.6044
21.4619 | -101.6209
21.4656 | -101.6386
21.4694 | -101.6581
21.4731 | -101.6784
21.4769 | -101.6994
21.4806 | -101.7225
21.4844 | -101.7469
21.4881 | -101.7697
21.4919 | -101.7913
21.4956 | -101.8123
21.4994 | -101.8307
21.5031 | -101.8501
21.5069 | -101.8731
21.5106 | -101.8984
21.5144 | -101.9224
21.5181 | -101.9431
21.5219 | -101.9615
21.5256 | -101.981
21.5294 | -102.0055
21.5331 | -102.028
21.5369 | -102.0468
21.5406 | -102.0627
21.5444 | -102.0777
21.5481 | -102.0915
21.5519 | -102.1072
21.5556 | -102.1267
21.5594 | -102.1444
21.5631 | -102.1589
21.5669 | -102.17
21.5706 | -102.1784
21.5744 | -102.1844
21.5781 | -102.1933
21.5819 | -102.2034
21.5856 | -102.2117
21.5894 | -102.2154
21.5931 | -102.2164
21.5969 | -102.2154
21.6006 | -102.2125
21.6044 | -102.2119
21.6081 | -102.2147
21.6119 | -102.2167
21.6156 | -102.2154
21.6194 | -102.2122
21.6231 | -102.2093
21.6269 | -102.2117
21.6306 | -102.2176
21.6344 | -102.2205
21.6381 | -102.2198
21.6419 | -102.2197
21.6456 | -102.2161
21.6494 | -102.2119
21.6531 | -102.2092
21.6569 | -102.2076
21.6607 | -102.2031
21.6644 | -102.1951
21.6682 | -102.1893
21.6719 | -102.1829
21.6757 | -102.1786
21.6794 | -102.1745
21.6832 | -102.1675
21.6869 | -102.1597
21.6907 | -102.1514
21.6944 | -102.1434
21.6982 | -102.137
21.7019 | -102.1315
21.7057 | -102.1256
21.7094 | -102.1189
21.7132 | -102.1108
21.7169 | -102.1009
21.7207 | -102.0905
21.7244 | -102.0796
21.7282 | -102.0696
21.7319 | -102.0576
21.7357 | -102.0443
21.7394 | -102.03
21.7432 | -102.0156
21.7469 | -102.0001
21.7507 | -101.9845
21.7544 | -101.9716
21.7582 | -101.9578
21.7619 | -101.9421
21.7657 | -101.9265
21.7694 | -101.9097
21.7732 | -101.8931
21.7769 | -101.8775
21.7807 | -101.8639
21.7844 | -101.8503
21.7882 | -101.8353
21.7919 | -101.8223
21.7957 | -101.8131
21.7994 | -101.8058
21.8032 | -101.7995
21.8069 | -101.7956
21.8107 | -101.7901
21.8144 | -101.7817
21.8182 | -101.7725
21.8219 | -101.767
21.8257 | -101.7639
21.8294 | -101.7627
21.8332 | -101.7627
21.8369 | -101.7619
21.8407 | -101.7627
21.8444 | -101.7638
21.8482 | -101.766
21.8519 | -101.7694
21.8557 | -101.7739
21.8594 | -101.7777
21.8632 | -101.7804
21.8669 | -101.7837
21.8707 | -101.7898
21.8744 | -101.7975
21.8782 | -101.8027
21.8819 | -101.8091
21.8857 | -101.8163
21.8894 | -101.8216
21.8932 | -101.8275
21.8969 | -101.8336
21.9007 | -101.8389
21.9044 | -101.8437
21.9082 | -101.8504
21.912 | -101.853
21.9157 | -101.854
21.9195 | -101.8567
21.9232 | -101.8584
21.927 | -101.8588
21.9307 | -101.8589
21.9345 | -101.8567
21.9382 | -101.8549
21.942 | -101.8533
21.9457 | -101.8527
21.9495 | -101.8518
21.9532 | -101.8527
21.957 | -101.8511
21.9607 | -101.8504
21.9645 | -101.8509
21.9682 | -101.8513
21.972 | -101.8533
21.9757 | -101.855
21.9795 | -101.855
21.9832 | -101.8545
21.987 | -101.8523
21.9907 | -101.851
21.9945 | -101.8518
21.9982 | -101.8533
22.002 | -101.8555
22.0057 | -101.8568
22.0095 | -101.8562
22.0132 | -101.8548
22.017 | -101.853
22.0207 | -101.8497
22.0245 | -101.8484
22.0282 | -101.8481
22.032 | -101.8482
22.0357 | -101.8484
22.0395 | -101.8488
22.0432 | -101.8482
22.047 | -101.8481
22.0507 | -101.8479
22.0545 | -101.8468
22.0582 | -101.8453
22.062 | -101.8441
22.0657 | -101.839
22.0695 | -101.8325
22.0732 | -101.8268
22.077 | -101.8225
22.0807 | -101.8199
22.0845 | -101.8165
22.0882 | -101.8127
22.092 | -101.8089
22.0957 | -101.8033
22.0995 | -101.7968
22.1032 | -101.7897
22.107 | -101.7819
22.1107 | -101.775
22.1145 | -101.7684
22.1182 | -101.7608
22.122 | -101.7525
22.1257 | -101.7439
22.1295 | -101.7355
22.1332 | -101.7253
22.137 | -101.7152
22.1407 | -101.7045
22.1445 | -101.6918
22.1482 | -101.6801
22.152 | -101.6678
22.1558 | -101.6562
22.1595 | -101.6475
22.1633 | -101.6369
22.167 | -101.6258
22.1708 | -101.6178
22.1745 | -101.611
22.1783 | -101.6045
22.182 | -101.5975
22.1858 | -101.5983
22.1895 | -101.5924
22.1933 | -101.5886
22.197 | -101.5856
22.2008 | -101.5829
22.2045 | -101.5833
22.2083 | -101.5842
22.212 | -101.5857
22.2158 | -101.5872
22.2195 | -101.5885
22.2233 | -101.5893
22.227 | -101.5954
22.2308 | -101.6036
22.2345 | -101.6113
22.2383 | -101.6197
22.242 | -101.6299
22.2458 | -101.639
22.2495 | -101.6481
22.2533 | -101.6582
22.257 | -101.6706
22.2608 | -101.6822
22.2645 | -101.695
22.2683 | -101.71
22.272 | -101.7271
22.2758 | -101.7447
22.2795 | -101.7612
22.2833 | -101.7767
22.287 | -101.7934
22.2908 | -101.8124
22.2945 | -101.8312
22.2983 | -101.8493
22.302 | -101.8637
22.3058 | -101.8786
22.3095 | -101.8914
22.3133 | -101.9061
22.317 | -101.9185
22.3208 | -101.9297
22.3245 | -101.9419
22.3283 | -101.9556
22.332 | -101.9707
22.3358 | -101.9845
22.3395 | -101.9954
22.3433 | -102.0061
22.347 | -102.0153
22.3508 | -102.0235
22.3545 | -102.0305
22.3583 | -102.039
22.362 | -102.0467
22.3658 | -102.0506
22.3695 | -102.0529
22.3733 | -102.0563
22.377 | -102.0587
22.3808 | -102.0625
22.3845 | -102.0667
22.3883 | -102.0679
22.392 | -102.0663
22.3958 | -102.0638
22.3996 | -102.0611
22.4033 | -102.0574
22.4071 | -102.0532
22.4108 | -102.05
22.4146 | -102.0472
22.4183 | -102.0423
22.4221 | -102.0378
22.4258 | -102.0333
22.4296 | -102.0275
22.4333 | -102.0225
22.4371 | -102.0181
22.4408 | -102.0133
22.4446 | -102.0092
22.4483 | -102.0037
22.4521 | -101.9986
22.4558 | -101.9925
22.4596 | -101.9897
22.4633 | -101.9841
22.4671 | -101.97
22.4708 | -101.9539
22.4746 | -101.9395
22.4783 | -101.9254
22.4821 | -101.9099
22.4858 | -101.8913
22.4896 | -101.8698
22.4933 | -101.8494
22.4971 | -101.8227
22.5008 | -101.7913
22.5046 | -101.7549
22.5083 | -101.7252
22.5121 | -101.6993
22.5158 | -101.6728
22.5196 | -101.6426
22.5233 | -101.6093
22.5271 | -101.5744
22.5308 | -101.538
22.5346 | -101.5002
22.5383 | -101.4621
22.5421 | -101.4215
22.5458 | -101.3743
22.5496 | -101.3242
22.5533 | -101.2743
22.5571 | -101.2231
22.5608 | -101.1699
22.5646 | -101.1185
22.5683 | -101.0682
22.5721 | -101.0169
22.5758 | -100.9649
22.5796 | -100.9121
22.5833 | -100.8597
22.5871 | -100.8066
22.5908 | -100.7547
22.5946 | -100.7049
22.5983 | -100.6581
22.6021 | -100.6136
22.6058 | -100.5706
22.6096 | -100.5296
22.6133 | -100.4911
22.6171 | -100.4567
22.6208 | -100.4268
22.6246 | -100.3991
22.6283 | -100.373
22.6321 | -100.3495
22.6358 | -100.3308
22.6396 | -100.3175
22.6433 | -100.3071
22.6471 | -100.2982
22.6509 | -100.293
22.6546 | -100.2924
22.6584 | -100.2956
22.6621 | -100.3007
22.6659 | -100.3094
22.6696 | -100.3226
22.6734 | -100.3377
22.6771 | -100.3549
22.6809 | -100.3764
22.6846 | -100.4021
22.6884 | -100.43
22.6921 | -100.4592
22.6959 | -100.4903
22.6996 | -100.5221
22.7034 | -100.5549
22.7071 | -100.59
22.7109 | -100.626
22.7146 | -100.664
22.7184 | -100.7037
22.7221 | -100.7429
22.7259 | -100.7758
22.7296 | -100.8134
22.7334 | -100.851
22.7371 | -100.8869
22.7409 | -100.9235
22.7446 | -100.9583
22.7484 | -100.9902
22.7521 | -101.0213
22.7559 | -101.0519
22.7596 | -101.081
22.7634 | -101.1087
22.7671 | -101.1339
22.7709 | -101.1576
22.7746 | -101.1795
22.7784 | -101.201
22.7821 | -101.2209
22.7859 | -101.2396
22.7896 | -101.2618
22.7934 | -101.2824
22.7971 | -101.297
22.8009 | -101.3086
22.8046 | -101.3189
22.8084 | -101.3281
22.8121 | -101.3384
22.8159 | -101.3427
22.8196 | -101.3497
22.8234 | -101.3601
22.8271 | -101.3692
22.8309 | -101.369
22.8346 | -101.3646
22.8384 | -101.3591
22.8421 | -101.3556
22.8459 | -101.3512
22.8496 | -101.3458
22.8534 | -101.343
22.8571 | -101.3409
22.8609 | -101.339
22.8646 | -101.3367
22.8684 | -101.3356
22.8721 | -101.3365
22.8759 | -101.3371
22.8796 | -101.3398
22.8834 | -101.3428
22.8871 | -101.343
22.8909 | -101.3421
22.8946 | -101.343
22.8984 | -101.3426
22.9022 | -101.3443
22.9059 | -101.3452
22.9097 | -101.3461
22.9134 | -101.3471
22.9172 | -101.3483
22.9209 | -101.3488
22.9247 | -101.3507
22.9284 | -101.3542
22.9322 | -101.3577
22.9359 | -101.3618
22.9397 | -101.3664
22.9434 | -101.3705
22.9472 | -101.3727
22.9509 | -101.3782
22.9547 | -101.3823
22.9584 | -101.3849
22.9622 | -101.3854
22.9659 | -101.3859
22.9697 | -101.3868
22.9734 | -101.3872
22.9772 | -101.3885
22.9809 | -101.3884
22.9847 | -101.3897
22.9884 | -101.3894
22.9922 | -101.3865
22.9959 | -101.3856
22.9997 | -101.3846
23.0034 | -101.3831
23.0072 | -101.3819
23.0109 | -101.3816
23.0147 | -101.3811
23.0184 | -101.3777
23.0222 | -101.3734
23.0259 | -101.3717
23.0297 | -101.3699
23.0334 | -101.3648
23.0372 | -101.3567
23.0409 | -101.3472
23.0447 | -101.339
23.0484 | -101.3324
23.0522 | -101.324
23.0559 | -101.3154
23.0597 | -101.3088
23.0634 | -101.3009
23.0672 | -101.289
23.0709 | -101.2795
23.0747 | -101.2719
23.0784 | -101.2642
23.0822 | -101.2542
23.0859 | -101.2427
23.0897 | -101.2279
23.0934 | -101.2129
23.0972 | -101.199
23.1009 | -101.1827
23.1047 | -101.1659
23.1084 | -101.146
23.1122 | -101.1252
23.1159 | -101.1056
23.1197 | -101.0858
23.1234 | -101.067
23.1272 | -101.0488
23.1309 | -101.0279
23.1347 | -101.0066
23.1384 | -100.9826
23.1422 | -100.9613
23.146 | -100.9404
23.1497 | -100.9172
23.1535 | -100.894
23.1572 | -100.8724
23.161 | -100.8493
23.1647 | -100.8258
23.1685 | -100.8043
23.1722 | -100.7841
23.176 | -100.7637
23.1797 | -100.7426
23.1835 | -100.7238
23.1872 | -100.7052
23.191 | -100.6894
23.1947 | -100.675
23.1985 | -100.6607
23.2022 | -100.6489
23.206 | -100.6382
23.2097 | -100.6281
23.2135 | -100.6177
23.2172 | -100.6059
23.221 | -100.5935
23.2247 | -100.5832
23.2285 | -100.5735
23.2322 | -100.567
23.236 | -100.5631
23.2397 | -100.5587
23.2435 | -100.5541
23.2472 | -100.5531
23.251 | -100.5513
23.2547 | -100.5482
23.2585 | -100.5469
23.2622 | -100.5453
23.266 | -100.5485
23.2697 | -100.5471
23.2735 | -100.5473
23.2772 | -100.5488
23.281 | -100.5478
23.2847 | -100.5461
23.2885 | -100.5451
23.2922 | -100.5439
23.296 | -100.5405
23.2997 | -100.5367
23.3035 | -100.5316
23.3072 | -100.525
23.311 | -100.5199
23.3147 | -100.5152
23.3185 | -100.5109
23.3222 | -100.5061
23.326 | -100.5001
23.3297 | -100.495
23.3335 | -100.488
23.3372 | -100.48
23.341 | -100.4721
23.3447 | -100.4636
23.3485 | -100.4541
23.3522 | -100.4449
23.356 | -100.4346
23.3597 | -100.4246
23.3635 | -100.417
23.3672 | -100.4094
23.371 | -100.403
23.3747 | -100.3933
23.3785 | -100.3827
23.3822 | -100.3729
23.386 | -100.3619
23.3898 | -100.3496
23.3935 | -100.339
23.3973 | -100.3305
23.401 | -100.3212
23.4048 | -100.3109
23.4085 | -100.3005
23.4123 | -100.2927
23.416 | -100.2859
23.4198 | -100.2823
23.4235 | -100.277
23.4273 | -100.2698
23.431 | -100.2643
23.4348 | -100.2589
23.4385 | -100.2546
23.4423 | -100.2494
23.446 | -100.2417
23.4498 | -100.2344
23.4535 | -100.2254
23.4573 | -100.2171
23.461 | -100.2124
23.4648 | -100.2083
23.4685 | -100.2029
23.4723 | -100.1963
23.476 | -100.1901
23.4798 | -100.1853
23.4835 | -100.1828
23.4873 | -100.1798
23.491 | -100.1777
23.4948 | -100.1762
23.4985 | -100.1725
23.5023 | -100.1678
23.506 | -100.1641
23.5098 | -100.1608
23.5135 | -100.1574
23.5173 | -100.1555
23.521 | -100.1521
23.5248 | -100.1478
23.5285 | -100.1437
23.5323 | -100.1391
23.536 | -100.1346
23.5398 | -100.1326
23.5435 | -100.1304
23.5473 | -100.1285
23.551 | -100.1267
23.5548 | -100.1251
23.5585 | -100.1245
23.5623 | -100.1232
23.566 | -100.121
23.5698 | -100.1197
23.5735 | -100.1192
23.5773 | -100.1151
23.581 | -100.1139
23.5848 | -100.1165
23.5885 | -100.1181
23.5923 | -100.1206
23.596 | -100.1244
23.5998 | -100.1277
23.6035 | -100.1331
23.6073 | -100.1386
23.611 | -100.1452
23.6148 | -100.1551
23.6185 | -100.1654
23.6223 | -100.1753
23.626 | -100.1853
23.6298 | -100.1985
23.6336 | -100.2104
23.6373 | -100.2223
23.6411 | -100.2349
23.6448 | -100.2474
23.6486 | -100.259
23.6523 | -100.269
23.6561 | -100.28
23.6598 | -100.2935
23.6636 | -100.3046
23.6673 | -100.3143
23.6711 | -100.3248
23.6748 | -100.3341
23.6786 | -100.3435
23.6823 | -100.3575
23.6861 | -100.3697
23.6898 | -100.3817
23.6936 | -100.3925
23.6973 | -100.3998
23.7011 | -100.409
23.7048 | -100.4194
23.7086 | -100.4318
23.7123 | -100.4447
23.7161 | -100.4544
23.7198 | -100.4637
23.7236 | -100.473
23.7273 | -100.4835
23.7311 | -100.4943
23.7348 | -100.5044
23.7386 | -100.515
23.7423 | -100.5241
23.7461 | -100.5313
23.7498 | -100.5403
23.7536 | -100.5481
23.7573 | -100.5577
23.7611 | -100.5672
23.7648 | -100.5766
23.7686 | -100.5861
23.7723 | -100.5951
23.7761 | -100.6044
23.7798 | -100.6167
23.7836 | -100.6279
23.7873 | -100.6383
23.7911 | -100.6476
23.7948 | -100.6557
23.7986 | -100.6623
23.8023 | -100.6714
23.8061 | -100.6815
23.8098 | -100.6914
23.8136 | -100.6991
23.8173 | -100.708
23.8211 | -100.7163
23.8248 | -100.7241
23.8286 | -100.7316
23.8323 | -100.7399
23.8361 | -100.7507
23.8398 | -100.7627
23.8436 | -100.7757
23.8473 | -100.7889
23.8511 | -100.8003
23.8548 | -100.8109
23.8586 | -100.821
23.8623 | -100.8339
23.8661 | -100.8486
23.8698 | -100.8639
23.8736 | -100.8808
23.8773 | -100.8954
23.8811 | -100.9106
23.8849 | -100.9279
23.8886 | -100.9454
23.8924 | -100.9627
23.8961 | -100.9775
23.8999 | -100.9917
23.9036 | -101.004
23.9074 | -101.0172
23.9111 | -101.0325
23.9149 | -101.0497
23.9186 | -101.0658
23.9224 | -101.079
23.9261 | -101.0915
23.9299 | -101.1038
23.9336 | -101.1145
23.9374 | -101.126
23.9411 | -101.1345
23.9449 | -101.1417
23.9486 | -101.1484
23.9524 | -101.1532
23.9561 | -101.1559
23.9599 | -101.1591
23.9636 | -101.163
23.9674 | -101.1622
23.9711 | -101.1629
23.9749 | -101.1642
23.9786 | -101.1626
23.9824 | -101.1588
23.9861 | -101.1546
23.9899 | -101.1509
23.9936 | -101.1457
23.9974 | -101.137
24.0011 | -101.1299
24.0049 | -101.1231
24.0086 | -101.1129
24.0124 | -101.1048
24.0161 | -101.0974
24.0199 | -101.0869
24.0236 | -101.0774
24.0274 | -101.0658
24.0311 | -101.0532
24.0349 | -101.0443
24.0386 | -101.0321
24.0424 | -101.0182
24.0461 | -101.0048
24.0499 | -100.9892
24.0536 | -100.975
24.0574 | -100.9612
24.0611 | -100.948
24.0649 | -100.9351
24.0686 | -100.9226
24.0724 | -100.9099
24.0761 | -100.8984
24.0799 | -100.889
24.0836 | -100.8791
24.0874 | -100.8677
24.0911 | -100.856
24.0949 | -100.8431
24.0986 | -100.8344
24.1024 | -100.8271
24.1061 | -100.8186
24.1099 | -100.8096
24.1136 | -100.8002
24.1174 | -100.7878
24.1212 | -100.7772
24.1249 | -100.7687
24.1287 | -100.7619
24.1324 | -100.7532
24.1362 | -100.7447
24.1399 | -100.7365
24.1437 | -100.7294
24.1474 | -100.7204
24.1512 | -100.7121
24.1549 | -100.7041
24.1587 | -100.6943
24.1624 | -100.6856
24.1662 | -100.6791
24.1699 | -100.6725
24.1737 | -100.666
24.1774 | -100.66
24.1812 | -100.655
24.1849 | -100.6495
24.1887 | -100.644
24.1924 | -100.6407
24.1962 | -100.6381
24.1999 | -100.6339
24.2037 | -100.629
24.2074 | -100.6237
24.2112 | -100.6179
24.2149 | -100.6126
24.2187 | -100.6088
24.2224 | -100.6057
24.2262 | -100.6018
24.2299 | -100.5986
24.2337 | -100.5948
24.2374 | -100.5925
24.2412 | -100.5904
24.2449 | -100.5887
24.2487 | -100.5844
24.2524 | -100.5798
24.2562 | -100.5724
24.2599 | -100.5561
24.2637 | -100.5318
24.2674 | -100.5102
24.2712 | -100.4964
24.2749 | -100.4873
24.2787 | -100.4779
24.2824 | -100.4677
24.2862 | -100.4585
24.2899 | -100.4485
24.2937 | -100.4378
24.2974 | -100.4285
24.3012 | -100.417
24.3049 | -100.4033
24.3087 | -100.3902
24.3124 | -100.3764
24.3162 | -100.361
24.3199 | -100.3444
24.3237 | -100.3261
24.3274 | -100.3074
24.3312 | -100.2904
24.3349 | -100.2738
24.3387 | -100.2572
24.3424 | -100.2393
24.3462 | -100.2223
24.3499 | -100.2052
24.3537 | -100.1879
24.3574 | -100.1694
24.3612 | -100.1507
24.3649 | -100.1329
24.3687 | -100.1158
24.3725 | -100.0997
24.3762 | -100.086
24.38 | -100.0721
24.3837 | -100.0595
24.3875 | -100.0453
24.3912 | -100.0332
24.395 | -100.0248
24.3987 | -100.0194
24.4025 | -100.0155
24.4062 | -100.0122
24.41 | -100.0094
24.4137 | -100.0075
24.4175 | -100.0081
24.4212 | -100.0111
24.425 | -100.0159
24.4287 | -100.0231
24.4325 | -100.0315
24.4362 | -100.041
24.44 | -100.0537
24.4437 | -100.0665
24.4475 | -100.0797
24.4512 | -100.0968
24.455 | -100.1178
24.4587 | -100.1407
24.4625 | -100.166
24.4662 | -100.1924
24.47 | -100.2194
24.4737 | -100.2458
24.4775 | -100.2738
24.4812 | -100.3041
24.485 | -100.3348
24.4887 | -100.3637
24.4925 | -100.3949
24.4962 | -100.4248
24.5 | -100.455
24.5037 | -100.4868
24.5075 | -100.519
24.5112 | -100.5499
24.515 | -100.581
24.5187 | -100.613
24.5225 | -100.6432
24.5262 | -100.6717
24.53 | -100.7
24.5337 | -100.7287
24.5375 | -100.7561
24.5412 | -100.7831
24.545 | -100.8124
24.5487 | -100.8407
24.5525 | -100.8678
24.5562 | -100.8943
24.56 | -100.9178
24.5637 | -100.9422
24.5675 | -100.9656
24.5712 | -100.9881
24.575 | -101.0082
24.5787 | -101.02
24.5825 | -101.0144
24.5862 | -100.9957
24.59 | -100.9619
24.5937 | -100.9402
24.5975 | -100.9268
24.6012 | -100.9192
24.605 | -100.9168
24.6087 | -100.9198
24.6125 | -100.9135
24.6163 | -100.8923
24.62 | -100.8643
24.6238 | -100.8341
24.6275 | -100.8026
24.6313 | -100.7679
24.635 | -100.7309
24.6388 | -100.6915
24.6425 | -100.6505
24.6463 | -100.6081
24.65 | -100.5646
24.6538 | -100.5196
24.6575 | -100.4723
24.6613 | -100.4238
24.665 | -100.3762
24.6688 | -100.3266
24.6725 | -100.2764
24.6763 | -100.2238
24.68 | -100.1697
24.6838 | -100.1151
24.6875 | -100.0621
24.6913 | -100.0102
24.695 | -99.9591
24.6988 | -99.9094
24.7025 | -99.8629
24.7063 | -99.8172
24.71 | -99.7735
24.7138 | -99.7318
24.7175 | -99.693
24.7213 | -99.6557
24.725 | -99.6209
24.7288 | -99.5895
24.7325 | -99.5612
24.7363 | -99.5389
24.74 | -99.5196
24.7438 | -99.5036
24.7475 | -99.4906
24.7513 | -99.4795
24.755 | -99.4739
24.7588 | -99.4736
24.7625 | -99.4775
24.7663 | -99.4846
24.77 | -99.4959
24.7738 | -99.5119
24.7775 | -99.5332
24.7813 | -99.559
24.785 | -99.5896
24.7888 | -99.6242
24.7925 | -99.6619
24.7963 | -99.7032
24.8 | -99.7476
24.8038 | -99.7943
24.8075 | -99.8441
24.8113 | -99.896
24.815 | -99.9489
24.8188 | -100.0039
24.8225 | -100.0623
24.8263 | -100.1239
24.83 | -100.1867
24.8338 | -100.249
24.8375 | -100.3146
24.8413 | -100.378
24.845 | -100.4425
24.8488 | -100.5088
24.8525 | -100.571
24.8563 | -100.631
24.8601 | -100.6908
24.8638 | -100.7494
24.8676 | -100.8053
24.8713 | -100.8591
24.8751 | -100.9117
24.8788 | -100.9601
24.8826 | -101.0056
24.8863 | -101.0496
24.8901 | -101.0626
24.8938 | -101.1
24.8976 | -101.1341
24.9013 | -101.1659
24.9051 | -101.1956
24.9088 | -101.2228
24.9126 | -101.2481
24.9163 | -101.2696
24.9201 | -101.2906
24.9238 | -101.3086
24.9276 | -101.3231
24.9313 | -101.3415
24.9351 | -101.3731
24.9388 | -101.4105
24.9426 | -101.4555
24.9463 | -101.4791
24.9501 | -101.4832
24.9538 | -101.4802
24.9576 | -101.4766
24.9613 | -101.4715
24.9651 | -101.4676
24.9688 | -101.4645
24.9726 | -101.462
24.9763 | -101.4602
24.9801 | -101.4608
24.9838 | -101.4617
24.9876 | -101.4634
24.9913 | -101.4656
24.9951 | -101.4688
24.9988 | -101.4736
25.0026 | -101.4782
25.0063 | -101.4854
25.0101 | -101.4957
25.0138 | -101.5079
25.0176 | -101.5188
25.0213 | -101.5315
25.0251 | -101.5421
25.0288 | -101.5539
25.0326 | -101.5688
25.0363 | -101.5854
25.0401 | -101.6018
25.0438 | -101.6182
25.0476 | -101.636
25.0513 | -101.6546
25.0551 | -101.6727
25.0588 | -101.6915
25.0626 | -101.7093
25.0663 | -101.7261
25.0701 | -101.7424
25.0738 | -101.7613
25.0776 | -101.7796
25.0813 | -101.7967
25.0851 | -101.8145
25.0888 | -101.8304
25.0926 | -101.8459
25.0963 | -101.8581
25.1001 | -101.8727
25.1038 | -101.8877
25.1076 | -101.9012
25.1114 | -101.9118
25.1151 | -101.9214
25.1189 | -101.9309
25.1226 | -101.9403
25.1264 | -101.948
25.1301 | -101.9525
25.1339 | -101.9574
25.1376 | -101.9624
25.1414 | -101.9644
25.1451 | -101.9649
25.1489 | -101.9664
25.1526 | -101.9666
25.1564 | -101.9666
25.1601 | -101.9666
25.1639 | -101.9655
25.1676 | -101.9638
25.1714 | -101.9632
25.1751 | -101.9613
25.1789 | -101.9581
25.1826 | -101.953
25.1864 | -101.9484
25.1901 | -101.9424
25.1939 | -101.9399
25.1976 | -101.9343
25.2014 | -101.9258
25.2051 | -101.9205
25.2089 | -101.9134
25.2126 | -101.905
25.2164 | -101.8973
25.2201 | -101.8911
25.2239 | -101.8852
25.2276 | -101.8791
25.2314 | -101.8758
25.2351 | -101.8719
25.2389 | -101.8677
25.2426 | -101.8637
25.2464 | -101.8583
25.2501 | -101.8554
25.2539 | -101.8524
25.2576 | -101.8492
25.2614 | -101.8447
25.2651 | -101.8414
25.2689 | -101.8387
25.2726 | -101.8358
25.2764 | -101.8327
25.2801 | -101.8323
25.2839 | -101.8341
25.2876 | -101.8367
25.2914 | -101.839
25.2951 | -101.8404
25.2989 | -101.8415
25.3026 | -101.8434
25.3064 | -101.8462
25.3101 | -101.8477
25.3139 | -101.8501
25.3176 | -101.8523
25.3214 | -101.8566
25.3251 | -101.8612
25.3289 | -101.8668
25.3326 | -101.8728
25.3364 | -101.8782
25.3401 | -101.8839
25.3439 | -101.8902
25.3476 | -101.8977
25.3514 | -101.904
25.3551 | -101.9094
25.3589 | -101.9161
25.3627 | -101.9215
25.3664 | -101.9258
25.3702 | -101.93
25.3739 | -101.9355
25.3777 | -101.9423
25.3814 | -101.9501
25.3852 | -101.9579
25.3889 | -101.9648
25.3927 | -101.9696
25.3964 | -101.9753
25.4002 | -101.9813
25.4039 | -101.9857
25.4077 | -101.9916
25.4114 | -101.9972
25.4152 | -102.0019
25.4189 | -102.0082
25.4227 | -102.0148
25.4264 | -102.0216
25.4302 | -102.0276
25.4339 | -102.0331
25.4377 | -102.0396
25.4414 | -102.0478
25.4452 | -102.0544
25.4489 | -102.0609
25.4527 | -102.0667
25.4564 | -102.0748
25.4602 | -102.0797
25.4639 | -102.0848
25.4677 | -102.09
25.4714 | -102.0958
25.4752 | -102.103
25.4789 | -102.1089
25.4827 | -102.1123
25.4864 | -102.1162
25.4902 | -102.119
25.4939 | -102.1202
25.4977 | -102.1205
25.5014 | -102.1215
25.5052 | -102.1238
25.5089 | -102.1238
25.5127 | -102.1255
25.5164 | -102.1256
25.5202 | -102.1227
25.5239 | -102.1203
25.5277 | -102.1192
25.5314 | -102.1163
25.5352 | -102.1106
25.5389 | -102.1034
25.5427 | -102.0984
25.5464 | -102.0929
25.5502 | -102.0837
25.5539 | -102.0735
25.5577 | -102.0655
25.5614 | -102.0561
25.5652 | -102.0458
25.5689 | -102.0359
25.5727 | -102.0259
25.5764 | -102.0145
25.5802 | -102.0009
25.5839 | -101.9878
25.5877 | -101.974
25.5914 | -101.959
25.5952 | -101.9445
25.5989 | -101.9291
25.6027 | -101.9139
25.6065 | -101.8995
25.6102 | -101.8853
25.614 | -101.8714
25.6177 | -101.8569
25.6215 | -101.8414
25.6252 | -101.8231
25.629 | -101.8067
25.6327 | -101.7926
25.6365 | -101.7777
25.6402 | -101.7602
25.644 | -101.7421
25.6477 | -101.7236
25.6515 | -101.7059
25.6552 | -101.6895
25.659 | -101.6721
25.6627 | -101.6567
25.6665 | -101.6426
25.6702 | -101.6296
25.674 | -101.6169
25.6777 | -101.604
25.6815 | -101.5907
25.6852 | -101.5804
25.689 | -101.5694
25.6927 | -101.5574
25.6965 | -101.5445
25.7002 | -101.5344
25.704 | -101.5267
25.7077 | -101.5209
25.7115 | -101.5163
25.7152 | -101.5106
25.719 | -101.5032
25.7227 | -101.4972
25.7265 | -101.4923
25.7302 | -101.4877
25.734 | -101.4828
25.7377 | -101.4783
25.7415 | -101.475
25.7452 | -101.4697
25.749 | -101.4649
25.7527 | -101.4632
25.7565 | -101.4611
25.7602 | -101.4573
25.764 | -101.4547
25.7677 | -101.4506
25.7715 | -101.4469
25.7752 | -101.4446
25.779 | -101.4416
25.7827 | -101.4378
25.7865 | -101.4338
25.7902 | -101.4291
25.794 | -101.425
25.7977 | -101.4205
25.8015 | -101.4144
25.8052 | -101.4088
25.809 | -101.4032
25.8127 | -101.3962
25.8165 | -101.3912
25.8202 | -101.3877
25.824 | -101.3834
25.8277 | -101.3789
25.8315 | -101.3746
25.8352 | -101.3702
25.839 | -101.3644
25.8427 | -101.3589
25.8465 | -101.3521
25.8503 | -101.3451
25.854 | -101.3401
25.8578 | -101.3373
25.8615 | -101.3324
25.8653 | -101.326
25.869 | -101.3186
25.8728 | -101.3131
25.8765 | -101.3079
25.8803 | -101.3032
25.884 | -101.3006
25.8878 | -101.2982
25.8915 | -101.2937
25.8953 | -101.2901
25.899 | -101.2883
25.9028 | -101.2857
25.9065 | -101.2817
25.9103 | -101.2792
25.914 | -101.2775
25.9178 | -101.2761
25.9215 | -101.2759
25.9253 | -101.2757
25.929 | -101.2736
25.9328 | -101.273
25.9365 | -101.2701
25.9403 | -101.2673
25.944 | -101.2654
25.9478 | -101.2604
25.9515 | -101.2538
25.9553 | -101.2477
25.959 | -101.2397
25.9628 | -101.2315
25.9665 | -101.2268
25.9703 | -101.2159
25.974 | -101.2102
25.9778 | -101.2
25.9815 | -101.1892
25.9853 | -101.1771
25.989 | -101.1671
25.9928 | -101.1598
25.9965 | -101.1509
26.0003 | -101.1404
26.004 | -101.1292
26.0078 | -101.1179
26.0115 | -101.107
26.0153 | -101.0956
26.019 | -101.083
26.0228 | -101.0685
26.0265 | -101.0529
26.0303 | -101.0355
26.034 | -101.0179
26.0378 | -101.0001
26.0415 | -100.9832
26.0453 | -100.9668
26.049 | -100.949
26.0528 | -100.9301
26.0565 | -100.9111
26.0603 | -100.894
26.064 | -100.8804
26.0678 | -100.8656
26.0715 | -100.8485
26.0753 | -100.8313
26.079 | -100.8154
26.0828 | -100.8023
26.0865 | -100.7922
26.0903 | -100.7816
26.094 | -100.7713
26.0978 | -100.7627
26.1016 | -100.7564
26.1053 | -100.7503
26.1091 | -100.7436
26.1128 | -100.7396
26.1166 | -100.7374
26.1203 | -100.735
26.1241 | -100.7333
26.1278 | -100.7339
26.1316 | -100.736
26.1353 | -100.7399
26.1391 | -100.7449
26.1428 | -100.7524
26.1466 | -100.7617
26.1503 | -100.7732
26.1541 | -100.7866
26.1578 | -100.8016
26.1616 | -100.8181
26.1653 | -100.837
26.1691 | -100.8566
26.1728 | -100.8765
26.1766 | -100.898
26.1803 | -100.92
26.1841 | -100.9403
26.1878 | -100.9617
26.1916 | -100.9856
26.1953 | -101.0119
26.1991 | -101.0396
26.2028 | -101.0687
26.2066 | -101.0989
26.2103 | -101.1289
26.2141 | -101.1617
26.2178 | -101.1946
26.2216 | -101.2271
26.2253 | -101.2573
26.2291 | -101.2881
26.2328 | -101.3211
26.2366 | -101.3529
26.2403 | -101.381
26.2441 | -101.4058
26.2478 | -101.431
26.2516 | -101.4551
26.2553 | -101.4765
26.2591 | -101.4988
26.2628 | -101.5188
26.2666 | -101.538
26.2703 | -101.5567
26.2741 | -101.5735
26.2778 | -101.5868
26.2816 | -101.5998
26.2853 | -101.6104
26.2891 | -101.6194
26.2928 | -101.6275
26.2966 | -101.635
26.3003 | -101.6413
26.3041 | -101.649
26.3078 | -101.6538
26.3116 | -101.6547
26.3153 | -101.6562
26.3191 | -101.6582
26.3228 | -101.6583
26.3266 | -101.6563
26.3303 | -101.6519
26.3341 | -101.6424
26.3378 | -101.6331
26.3416 | -101.6227
26.3454 | -101.6112
26.3491 | -101.6025
26.3529 | -101.5913
26.3566 | -101.579
26.3604 | -101.5659
26.3641 | -101.5523
26.3679 | -101.5378
26.3716 | -101.5237
26.3754 | -101.5098
26.3791 | -101.4951
26.3829 | -101.4804
26.3866 | -101.4657
26.3904 | -101.4502
26.3941 | -101.4354
26.3979 | -101.4222
26.4016 | -101.4047
26.4054 | -101.3816
26.4091 | -101.3603
26.4129 | -101.34
26.4166 | -101.318
26.4204 | -101.2985
26.4241 | -101.2814
26.4279 | -101.2663
26.4316 | -101.2495
26.4354 | -101.2302
26.4391 | -101.213
26.4429 | -101.1974
26.4466 | -101.1813
26.4504 | -101.1649
26.4541 | -101.1471
26.4579 | -101.1295
26.4616 | -101.1123
26.4654 | -101.096
26.4691 | -101.0786
26.4729 | -101.0606
26.4766 | -101.0434
26.4804 | -101.0265
26.4841 | -101.0084
26.4879 | -100.9908
26.4916 | -100.9729
26.4954 | -100.9537
26.4991 | -100.9347
26.5029 | -100.9165
26.5066 | -100.8976
26.5104 | -100.8794
26.5141 | -100.8492
26.5179 | -100.8218
26.5216 | -100.7977
26.5254 | -100.7753
26.5291 | -100.7525
26.5329 | -100.7293
26.5366 | -100.704
26.5404 | -100.6791
26.5441 | -100.6572
26.5479 | -100.636
26.5516 | -100.6153
26.5554 | -100.5969
26.5591 | -100.5787
26.5629 | -100.5585
26.5666 | -100.5389
26.5704 | -100.5193
26.5741 | -100.4991
26.5779 | -100.4804
26.5816 | -100.4623
26.5854 | -100.4448
26.5892 | -100.4258
26.5929 | -100.4084
26.5967 | -100.3919
26.6004 | -100.3766
26.6042 | -100.3623
26.6079 | -100.347
26.6117 | -100.3331
26.6154 | -100.3215
26.6192 | -100.3113
26.6229 | -100.3008
26.6267 | -100.2905
26.6304 | -100.2808
26.6342 | -100.2689
26.6379 | -100.2591
26.6417 | -100.2506
26.6454 | -100.2412
26.6492 | -100.2332
26.6529 | -100.2277
26.6567 | -100.2197
26.6604 | -100.2112
26.6642 | -100.2055
26.6679 | -100.2007
26.6717 | -100.1954
26.6754 | -100.1879
26.6792 | -100.1801
26.6829 | -100.1749
26.6867 | -100.1703
26.6904 | -100.1664
26.6942 | -100.1629
26.6979 | -100.1599
26.7017 | -100.1572
26.7054 | -100.1545
26.7092 | -100.1503
26.7129 | -100.1471
26.7167 | -100.1447
26.7204 | -100.1417
26.7242 | -100.1377
26.7279 | -100.1347
26.7317 | -100.1311
26.7354 | -100.1232
26.7392 | -100.116
26.7429 | -100.1106
26.7467 | -100.1019
26.7504 | -100.0924
26.7542 | -100.0818
26.7579 | -100.0704
26.7617 | -100.0605
26.7654 | -100.0507
26.7692 | -100.0406
26.7729 | -100.0281
26.7767 | -100.0158
26.7804 | -100.004
26.7842 | -99.9899
26.7879 | -99.9755
26.7917 | -99.9614
26.7954 | -99.9465
26.7992 | -99.9275
26.8029 | -99.9076
26.8067 | -99.887
26.8104 | -99.8653
26.8142 | -99.8445
26.8179 | -99.8233
26.8217 | -99.8008
26.8254 | -99.7773
26.8292 | -99.7542
26.833 | -99.7316
26.8367 | -99.7079
26.8405 | -99.6837
26.8442 | -99.6615
26.848 | -99.6391
26.8517 | -99.6149
26.8555 | -99.5909
26.8592 | -99.5679
26.863 | -99.5449
26.8667 | -99.522
26.8705 | -99.4995
26.8742 | -99.4766
26.878 | -99.4529
26.8817 | -99.4314
26.8855 | -99.4097
26.8892 | -99.389
26.893 | -99.372
26.8967 | -99.3542
26.9005 | -99.3356
26.9042 | -99.3154
26.908 | -99.294
26.9117 | -99.2724
26.9155 | -99.2526
26.9192 | -99.232
26.923 | -99.2111
26.9267 | -99.1893
26.9305 | -99.1695
26.9342 | -99.149
26.938 | -99.1286
26.9417 | -99.109
26.9455 | -99.0883
26.9492 | -99.0686
26.953 | -99.0501
26.9567 | -99.0324
26.9605 | -99.0156
26.9642 | -98.9992
26.968 | -98.9824
26.9717 | -98.964
26.9755 | -98.9459
26.9792 | -98.9284
26.983 | -98.9092
26.9867 | -98.8937
26.9905 | -98.8771
26.9942 | -98.8576
26.998 | -98.8363
27.0017 | -98.8175
27.0055 | -98.8018
27.0092 | -98.7867
27.013 | -98.7726
27.0167 | -98.7588
27.0205 | -98.7423
27.0242 | -98.724
27.028 | -98.7078
27.0317 | -98.6927
27.0355 | -98.6776
27.0392 | -98.6635
27.043 | -98.6472
27.0467 | -98.6304
27.0505 | -98.6135
27.0542 | -98.5944
27.058 | -98.5801
27.0617 | -98.5647
27.0655 | -98.5495
27.0692 | -98.5342
27.073 | -98.5191
27.0768 | -98.5027
27.0805 | -98.4828
27.0843 | -98.4566
27.088 | -98.4309
27.0918 | -98.4051
27.0955 | -98.3785
27.0993 | -98.3558
27.103 | -98.3311
27.1068 | -98.3069
27.1105 | -98.2824
27.1143 | -98.2598
27.118 | -98.2387
27.1218 | -98.2167
27.1255 | -98.1931
27.1293 | -98.1706
27.133 | -98.1484
27.1368 | -98.1262
27.1405 | -98.1019
27.1443 | -98.0754
27.148 | -98.0494
27.1518 | -98.0253
27.1555 | -98.0023
27.1593 | -97.9789
27.163 | -97.956
27.1668 | -97.9339
27.1705 | -97.9103
27.1743 | -97.8875
27.178 | -97.8648
27.1818 | -97.843
27.1855 | -97.8214
27.1893 | -97.8022
27.193 | -97.7839
27.1968 | -97.7641
27.2005 | -97.7453
27.2043 | -97.7282
27.208 | -97.7121
27.2118 | -97.6977
27.2155 | -97.6832
27.2193 | -97.67
27.223 | -97.6588
27.2268 | -97.6481
27.2305 | -97.639
27.2343 | -97.6319
27.238 | -97.6267
27.2418 | -97.6222
27.2455 | -97.6176
27.2493 | -97.6151
27.253 | -97.6146
27.2568 | -97.6161
27.2605 | -97.6178
27.2643 | -97.619
27.268 | -97.6226
27.2718 | -97.6298
27.2755 | -97.6372
27.2793 | -97.6477
27.283 | -97.6583
27.2868 | -97.6711
27.2905 | -97.6856
27.2943 | -97.7032
27.298 | -97.7228
27.3018 | -97.7434
27.3055 | -97.7645
27.3093 | -97.7883
27.313 | -97.8116
27.3168 | -97.8374
27.3205 | -97.8648
27.3243 | -97.8931
27.3281 | -97.9234
27.3318 | -97.9544
27.3356 | -97.9857
27.3393 | -98.0178
27.3431 | -98.0504
27.3468 | -98.0842
27.3506 | -98.1193
27.3543 | -98.1554
27.3581 | -98.1909
27.3618 | -98.2253
27.3656 | -98.262
27.3693 | -98.2955
27.3731 | -98.3308
27.3768 | -98.3695
27.3806 | -98.4078
27.3843 | -98.4429
27.3881 | -98.4754
27.3918 | -98.5083
27.3956 | -98.541
27.3993 | -98.5715
27.4031 | -98.602
27.4068 | -98.6343
27.4106 | -98.6638
27.4143 | -98.6907
27.4181 | -98.7183
27.4218 | -98.7429
27.4256 | -98.7663
27.4293 | -98.7888
27.4331 | -98.8111
27.4368 | -98.8344
27.4406 | -98.8535
27.4443 | -98.8723
27.4481 | -98.8911
27.4518 | -98.9082
27.4556 | -98.9248
27.4593 | -98.9388
27.4631 | -98.9533
27.4668 | -98.9702
27.4706 | -98.9934
27.4743 | -99.014
27.4781 | -99.0297
27.4818 | -99.0428
27.4856 | -99.0515
27.4893 | -99.0608
27.4931 | -99.0686
27.4968 | -99.0772
27.5006 | -99.0817
27.5043 | -99.0834
27.5081 | -99.0848
27.5118 | -99.0881
27.5156 | -99.0915
27.5193 | -99.0952
27.5231 | -99.0999
27.5268 | -99.1057
27.5306 | -99.1114
27.5343 | -99.1133
27.5381 | -99.1086
27.5418 | -99.1011
27.5456 | -99.0938
27.5493 | -99.0819
27.5531 | -99.0659
27.5568 | -99.0519
27.5606 | -99.0421
27.5643 | -99.0336
27.5681 | -99.0255
27.5719 | -99.0183
27.5756 | -99.0083
27.5794 | -98.9977
27.5831 | -98.9874
27.5869 | -98.9767
27.5906 | -98.9667
27.5944 | -98.9445
27.5981 | -98.9305
27.6019 | -98.9146
27.6056 | -98.9
27.6094 | -98.8851
27.6131 | -98.8682
27.6169 | -98.8513
27.6206 | -98.8321
27.6244 | -98.8137
27.6281 | -98.7943
27.6319 | -98.7756
27.6356 | -98.7557
27.6394 | -98.7351
27.6431 | -98.7131
27.6469 | -98.6907
27.6506 | -98.6692
27.6544 | -98.6486
27.6581 | -98.626
27.6619 | -98.6028
27.6656 | -98.5817
27.6694 | -98.5594
27.6731 | -98.5368
27.6769 | -98.5119
27.6806 | -98.4905
27.6844 | -98.4705
27.6881 | -98.4484
27.6919 | -98.4276
27.6956 | -98.4067
27.6994 | -98.3867
27.7031 | -98.368
27.7069 | -98.3483
27.7106 | -98.3322
27.7144 | -98.3185
27.7181 | -98.305
27.7219 | -98.2923
27.7256 | -98.281
27.7294 | -98.2715
27.7331 | -98.2631
27.7369 | -98.255
27.7406 | -98.2486
27.7444 | -98.2436
27.7481 | -98.2417
27.7519 | -98.2415
27.7556 | -98.242
27.7594 | -98.2447
27.7631 | -98.2467
27.7669 | -98.2482
27.7706 | -98.2518
27.7744 | -98.2575
27.7781 | -98.2652
27.7819 | -98.2732
27.7856 | -98.2825
27.7894 | -98.2936
27.7931 | -98.307
27.7969 | -98.3214
27.8006 | -98.3353
27.8044 | -98.3512
27.8081 | -98.3679
27.8119 | -98.3863
27.8157 | -98.4055
27.8194 | -98.4266
27.8232 | -98.4492
27.8269 | -98.4723
27.8307 | -98.4948
27.8344 | -98.5157
27.8382 | -98.5385
27.8419 | -98.561
27.8457 | -98.5841
27.8494 | -98.6097
27.8532 | -98.6338
27.8569 | -98.6571
27.8607 | -98.6809
27.8644 | -98.7064
27.8682 | -98.7352
27.8719 | -98.763
27.8757 | -98.7897
27.8794 | -98.8177
27.8832 | -98.8462
27.8869 | -98.8717
27.8907 | -98.8981
27.8944 | -98.9273
27.8982 | -98.9566
27.9019 | -98.9842
27.9057 | -99.0109
27.9094 | -99.0365
27.9132 | -99.0622
27.9169 | -99.088
27.9207 | -99.1126
27.9244 | -99.1401
27.9282 | -99.169
27.9319 | -99.1983
27.9357 | -99.2284
27.9394 | -99.2608
27.9432 | -99.2934
27.9469 | -99.3217
27.9507 | -99.3431
27.9544 | -99.3606
27.9582 | -99.3775
27.9619 | -99.3916
27.9657 | -99.4061
27.9694 | -99.4218
27.9732 | -99.4382
27.9769 | -99.4508
27.9807 | -99.4601
27.9844 | -99.4726
27.9882 | -99.4843
27.9919 | -99.4944
27.9957 | -99.5036
27.9994 | -99.51
28.0032 | -99.5132