## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 252 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 270
Hood | 02-92B/P | 194
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1017.7 | hPa |  | PRC PRO 01
Humidity | 57.5 | %/Hr
Internal temperature of calibration bench | 23.77 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1027 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.9 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.97 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.99 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3105 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 19/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 25/06/2018
Technical Manager | Anthony HUE | Approval |  | 28/06/2018
CEO | Robert GLEMOT | Approval |  | 28/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5242 | 57.4735
0.012209 | -36.0431 | 38.3778
0.018313 | -34.9835 | 28.4756
0.024417 | -34.5338 | 22.1224
0.030521 | -34.4332 | 18.3595
0.036626 | -34.2919 | 15.2533
0.04273 | -34.2089 | 12.9174
0.048834 | -34.1579 | 11.3479
0.054938 | -34.0841 | 10.1517
0.061043 | -34.0047 | 9.007
0.067147 | -33.9948 | 8.1312
0.073251 | -33.9975 | 7.3738
0.079355 | -34.0184 | 6.8133
0.08546 | -34.0383 | 6.3506
0.091564 | -34.0348 | 5.9786
0.097668 | -33.9991 | 5.7398
0.10377 | -33.9493 | 5.4005
0.10988 | -33.9169 | 5.0405
0.11598 | -33.9074 | 4.6777
0.12209 | -33.9251 | 4.3944
0.12819 | -33.9597 | 4.156
0.13429 | -33.9933 | 3.9088
0.1404 | -34.0113 | 3.6267
0.1465 | -34.0276 | 3.3606
0.15261 | -34.0347 | 3.1143
0.15871 | -34.0306 | 2.9148
0.16482 | -34.0152 | 2.7578
0.17092 | -33.99 | 2.7043
0.17702 | -33.9706 | 2.612
0.18313 | -33.9529 | 2.534
0.18923 | -33.9403 | 2.4298
0.19534 | -33.9334 | 2.3261
0.20144 | -33.9325 | 2.2105
0.20754 | -33.9352 | 2.1344
0.21365 | -33.9377 | 2.0581
0.21975 | -33.942 | 1.9924
0.22586 | -33.9501 | 1.9265
0.23196 | -33.9588 | 1.8537
0.23807 | -33.9677 | 1.7763
0.24417 | -33.97 | 1.6829
0.25027 | -33.9736 | 1.5986
0.25638 | -33.9736 | 1.5062
0.26248 | -33.9717 | 1.4185
0.26859 | -33.9694 | 1.3345
0.27469 | -33.9687 | 1.2466
0.2808 | -33.9748 | 1.1805
0.2869 | -33.9794 | 1.1106
0.293 | -33.9851 | 1.0522
0.29911 | -33.9916 | 1.0131
0.30521 | -33.9978 | 0.97596
0.31132 | -34.0019 | 0.9632
0.31742 | -34.0063 | 0.94178
0.32353 | -34.0098 | 0.925
0.32963 | -34.0127 | 0.91179
0.33573 | -34.0147 | 0.90118
0.34184 | -34.0173 | 0.88711
0.34794 | -34.0206 | 0.8538
0.35405 | -34.024 | 0.82573
0.36015 | -34.0277 | 0.79399
0.36626 | -34.0302 | 0.75837
0.37236 | -34.0327 | 0.70907
0.37846 | -34.0351 | 0.65156
0.38457 | -34.0341 | 0.59085
0.39067 | -34.0348 | 0.53997
0.39678 | -34.0332 | 0.48236
0.40288 | -34.0308 | 0.42759
0.40899 | -34.027 | 0.37413
0.41509 | -34.0224 | 0.32024
0.42119 | -34.0173 | 0.27129
0.4273 | -34.0135 | 0.2206
0.4334 | -34.01 | 0.16903
0.43951 | -34.0069 | 0.1184
0.44561 | -34.0054 | 0.071165
0.45172 | -34.007 | 0.032476
0.45782 | -34.0067 | -0.010897
0.46392 | -34.008 | -0.050725
0.47003 | -34.0097 | -0.086791
0.47613 | -34.0118 | -0.11982
0.48224 | -34.014 | -0.15133
0.48834 | -34.0163 | -0.17364
0.49445 | -34.0183 | -0.19943
0.50055 | -34.0208 | -0.22251
0.50665 | -34.0225 | -0.24157
0.51276 | -34.0236 | -0.26219
0.51886 | -34.0247 | -0.28026
0.52497 | -34.0237 | -0.2866
0.53107 | -34.0241 | -0.29716
0.53717 | -34.0244 | -0.30857
0.54328 | -34.0242 | -0.3134
0.54938 | -34.0242 | -0.32122
0.55549 | -34.0237 | -0.32979
0.56159 | -34.0237 | -0.33983
0.5677 | -34.0233 | -0.34943
0.5738 | -34.0229 | -0.36206
0.5799 | -34.023 | -0.37344
0.58601 | -34.0235 | -0.38895
0.59211 | -34.0247 | -0.41053
0.59822 | -34.0255 | -0.42631
0.60432 | -34.0267 | -0.44728
0.61043 | -34.0278 | -0.46769
0.61653 | -34.0291 | -0.48861
0.62263 | -34.0303 | -0.51101
0.62874 | -34.0317 | -0.5308
0.63484 | -34.0329 | -0.55136
0.64095 | -34.0342 | -0.5718
0.64705 | -34.0354 | -0.59112
0.65316 | -34.0365 | -0.61013
0.65926 | -34.0375 | -0.63243
0.66536 | -34.0369 | -0.65063
0.67147 | -34.0371 | -0.66636
0.67757 | -34.0377 | -0.68388
0.68368 | -34.0374 | -0.70377
0.68978 | -34.0375 | -0.72155
0.69589 | -34.037 | -0.73972
0.70199 | -34.0361 | -0.76179
0.70809 | -34.0355 | -0.78188
0.7142 | -34.0349 | -0.80121
0.7203 | -34.034 | -0.82072
0.72641 | -34.0331 | -0.8418
0.73251 | -34.0328 | -0.86266
0.73862 | -34.0328 | -0.88741
0.74472 | -34.0331 | -0.90689
0.75082 | -34.0326 | -0.92829
0.75693 | -34.0327 | -0.95215
0.76303 | -34.0327 | -0.97508
0.76914 | -34.0334 | -1.0027
0.77524 | -34.0337 | -1.028
0.78135 | -34.0342 | -1.0518
0.78745 | -34.0349 | -1.0771
0.79355 | -34.0355 | -1.1021
0.79966 | -34.0367 | -1.1285
0.80576 | -34.038 | -1.1596
0.81187 | -34.0384 | -1.1898
0.81797 | -34.0392 | -1.2213
0.82408 | -34.04 | -1.2492
0.83018 | -34.0407 | -1.2811
0.83628 | -34.0417 | -1.3122
0.84239 | -34.0416 | -1.341
0.84849 | -34.0417 | -1.3699
0.8546 | -34.0422 | -1.3983
0.8607 | -34.0423 | -1.4284
0.86681 | -34.0425 | -1.4574
0.87291 | -34.0421 | -1.4837
0.87901 | -34.0416 | -1.5065
0.88512 | -34.0414 | -1.5299
0.89122 | -34.0413 | -1.5538
0.89733 | -34.0411 | -1.5733
0.90343 | -34.0408 | -1.595
0.90953 | -34.041 | -1.6098
0.91564 | -34.0404 | -1.6259
0.92174 | -34.04 | -1.64
0.92785 | -34.0395 | -1.6535
0.93395 | -34.0394 | -1.6651
0.94006 | -34.0394 | -1.6785
0.94616 | -34.0392 | -1.6827
0.95226 | -34.0392 | -1.6939
0.95837 | -34.0393 | -1.7053
0.96447 | -34.0397 | -1.7152
0.97058 | -34.0397 | -1.7248
0.97668 | -34.0394 | -1.7324
0.98279 | -34.0394 | -1.7356
0.98889 | -34.0396 | -1.7397
0.99499 | -34.0396 | -1.7479
1.0011 | -34.0398 | -1.7522
1.0072 | -34.0396 | -1.7554
1.0133 | -34.0395 | -1.7586
1.0194 | -34.0395 | -1.7637
1.0255 | -34.0391 | -1.7702
1.0316 | -34.0384 | -1.7738
1.0377 | -34.0378 | -1.7793
1.0438 | -34.037 | -1.7834
1.0499 | -34.0353 | -1.7898
1.056 | -34.0341 | -1.7944
1.0621 | -34.0328 | -1.8
1.0682 | -34.0323 | -1.8046
1.0743 | -34.031 | -1.8124
1.0805 | -34.0303 | -1.8191
1.0866 | -34.0286 | -1.83
1.0927 | -34.0271 | -1.8374
1.0988 | -34.0262 | -1.8481
1.1049 | -34.0246 | -1.8556
1.111 | -34.0229 | -1.8668
1.1171 | -34.0214 | -1.8763
1.1232 | -34.02 | -1.8907
1.1293 | -34.0188 | -1.9026
1.1354 | -34.0177 | -1.9152
1.1415 | -34.0165 | -1.9295
1.1476 | -34.015 | -1.9436
1.1537 | -34.0135 | -1.9571
1.1598 | -34.0134 | -1.9756
1.1659 | -34.0128 | -1.9876
1.172 | -34.012 | -2.0009
1.1781 | -34.0112 | -2.015
1.1842 | -34.0106 | -2.0307
1.1903 | -34.0114 | -2.0457
1.1964 | -34.0112 | -2.0602
1.2025 | -34.0111 | -2.0756
1.2086 | -34.0116 | -2.0913
1.2147 | -34.0119 | -2.1062
1.2209 | -34.0126 | -2.1195
1.227 | -34.0146 | -2.1344
1.2331 | -34.0156 | -2.1485
1.2392 | -34.0166 | -2.1642
1.2453 | -34.0178 | -2.176
1.2514 | -34.0192 | -2.1879
1.2575 | -34.0208 | -2.2017
1.2636 | -34.0233 | -2.2128
1.2697 | -34.0251 | -2.2241
1.2758 | -34.0267 | -2.2381
1.2819 | -34.0286 | -2.2513
1.288 | -34.0304 | -2.2648
1.2941 | -34.0325 | -2.275
1.3002 | -34.0346 | -2.2828
1.3063 | -34.0364 | -2.2924
1.3124 | -34.0383 | -2.302
1.3185 | -34.0404 | -2.3123
1.3246 | -34.0423 | -2.3227
1.3307 | -34.0445 | -2.3317
1.3368 | -34.0465 | -2.3388
1.3429 | -34.0482 | -2.3471
1.349 | -34.0497 | -2.3552
1.3551 | -34.0515 | -2.3658
1.3613 | -34.0531 | -2.3731
1.3674 | -34.0545 | -2.3813
1.3735 | -34.0558 | -2.3889
1.3796 | -34.0572 | -2.3959
1.3857 | -34.0582 | -2.4024
1.3918 | -34.0593 | -2.4119
1.3979 | -34.0607 | -2.42
1.404 | -34.0614 | -2.4298
1.4101 | -34.0624 | -2.4385
1.4162 | -34.0633 | -2.4467
1.4223 | -34.0641 | -2.4569
1.4284 | -34.0649 | -2.465
1.4345 | -34.0656 | -2.4751
1.4406 | -34.0661 | -2.4863
1.4467 | -34.0668 | -2.4955
1.4528 | -34.0675 | -2.5046
1.4589 | -34.0683 | -2.515
1.465 | -34.0686 | -2.5225
1.4711 | -34.0695 | -2.5314
1.4772 | -34.0701 | -2.5436
1.4833 | -34.0707 | -2.556
1.4894 | -34.0715 | -2.5674
1.4955 | -34.0723 | -2.5789
1.5016 | -34.0729 | -2.591
1.5078 | -34.0734 | -2.6059
1.5139 | -34.0741 | -2.6186
1.52 | -34.0746 | -2.6319
1.5261 | -34.075 | -2.6436
1.5322 | -34.0752 | -2.6564
1.5383 | -34.0755 | -2.6687
1.5444 | -34.0762 | -2.6822
1.5505 | -34.0766 | -2.6957
1.5566 | -34.0772 | -2.7087
1.5627 | -34.0778 | -2.7222
1.5688 | -34.0786 | -2.7356
1.5749 | -34.0789 | -2.7481
1.581 | -34.0792 | -2.7651
1.5871 | -34.0798 | -2.781
1.5932 | -34.0802 | -2.794
1.5993 | -34.0806 | -2.8112
1.6054 | -34.081 | -2.8264
1.6115 | -34.0813 | -2.8437
1.6176 | -34.0817 | -2.8602
1.6237 | -34.0819 | -2.875
1.6298 | -34.0822 | -2.8906
1.6359 | -34.0826 | -2.9053
1.642 | -34.0831 | -2.92
1.6482 | -34.0835 | -2.9372
1.6543 | -34.084 | -2.9543
1.6604 | -34.0844 | -2.969
1.6665 | -34.0845 | -2.9842
1.6726 | -34.0848 | -3.0006
1.6787 | -34.0853 | -3.0179
1.6848 | -34.0851 | -3.0345
1.6909 | -34.0855 | -3.0509
1.697 | -34.0857 | -3.0673
1.7031 | -34.0858 | -3.0831
1.7092 | -34.0857 | -3.0982
1.7153 | -34.0859 | -3.1145
1.7214 | -34.0855 | -3.1299
1.7275 | -34.0853 | -3.1459
1.7336 | -34.085 | -3.1623
1.7397 | -34.0851 | -3.1778
1.7458 | -34.0851 | -3.1945
1.7519 | -34.0849 | -3.2108
1.758 | -34.0848 | -3.2257
1.7641 | -34.0849 | -3.2402
1.7702 | -34.0846 | -3.2562
1.7763 | -34.0846 | -3.2699
1.7824 | -34.0844 | -3.2862
1.7885 | -34.0842 | -3.3003
1.7947 | -34.0843 | -3.3152
1.8008 | -34.0842 | -3.3281
1.8069 | -34.0841 | -3.3417
1.813 | -34.0841 | -3.356
1.8191 | -34.0838 | -3.3704
1.8252 | -34.0835 | -3.3844
1.8313 | -34.0836 | -3.3989
1.8374 | -34.0836 | -3.4124
1.8435 | -34.0836 | -3.4246
1.8496 | -34.0833 | -3.4392
1.8557 | -34.0832 | -3.4541
1.8618 | -34.0832 | -3.4659
1.8679 | -34.0832 | -3.479
1.874 | -34.0832 | -3.4922
1.8801 | -34.0831 | -3.5052
1.8862 | -34.0832 | -3.518
1.8923 | -34.0831 | -3.5302
1.8984 | -34.0828 | -3.5409
1.9045 | -34.0827 | -3.5519
1.9106 | -34.0824 | -3.5638
1.9167 | -34.0821 | -3.5751
1.9228 | -34.0821 | -3.5859
1.9289 | -34.0817 | -3.5948
1.9351 | -34.0818 | -3.6038
1.9412 | -34.0819 | -3.614
1.9473 | -34.082 | -3.626
1.9534 | -34.0819 | -3.6371
1.9595 | -34.0816 | -3.6486
1.9656 | -34.0812 | -3.6594
1.9717 | -34.0809 | -3.6699
1.9778 | -34.081 | -3.6815
1.9839 | -34.081 | -3.6929
1.99 | -34.081 | -3.7027
1.9961 | -34.0808 | -3.7131
2.0022 | -34.0807 | -3.7231
2.0083 | -34.0803 | -3.733
2.0144 | -34.0803 | -3.7428
2.0205 | -34.0804 | -3.7547
2.0266 | -34.0805 | -3.7663
2.0327 | -34.0805 | -3.7774
2.0388 | -34.0808 | -3.7914
2.0449 | -34.0808 | -3.8016
2.051 | -34.0809 | -3.8116
2.0571 | -34.0807 | -3.823
2.0632 | -34.0806 | -3.8336
2.0693 | -34.0806 | -3.8448
2.0754 | -34.0808 | -3.8567
2.0816 | -34.081 | -3.8687
2.0877 | -34.0809 | -3.8798
2.0938 | -34.0811 | -3.8923
2.0999 | -34.0811 | -3.904
2.106 | -34.0813 | -3.916
2.1121 | -34.0813 | -3.9284
2.1182 | -34.0814 | -3.9384
2.1243 | -34.0816 | -3.9489
2.1304 | -34.0817 | -3.9592
2.1365 | -34.0821 | -3.9708
2.1426 | -34.0822 | -3.9842
2.1487 | -34.0826 | -3.996
2.1548 | -34.0831 | -4.0091
2.1609 | -34.0836 | -4.0212
2.167 | -34.0841 | -4.0338
2.1731 | -34.0844 | -4.0459
2.1792 | -34.0851 | -4.0571
2.1853 | -34.0857 | -4.0701
2.1914 | -34.0861 | -4.0824
2.1975 | -34.0867 | -4.0946
2.2036 | -34.0872 | -4.1079
2.2097 | -34.0876 | -4.12
2.2158 | -34.0886 | -4.1303
2.222 | -34.0891 | -4.1403
2.2281 | -34.0894 | -4.1512
2.2342 | -34.0899 | -4.1644
2.2403 | -34.0904 | -4.1744
2.2464 | -34.0912 | -4.1856
2.2525 | -34.0916 | -4.1964
2.2586 | -34.0925 | -4.2062
2.2647 | -34.0934 | -4.2155
2.2708 | -34.094 | -4.2272
2.2769 | -34.0948 | -4.2387
2.283 | -34.0953 | -4.2482
2.2891 | -34.0959 | -4.2585
2.2952 | -34.0964 | -4.2683
2.3013 | -34.0971 | -4.2785
2.3074 | -34.0978 | -4.2881
2.3135 | -34.0983 | -4.2995
2.3196 | -34.0989 | -4.308
2.3257 | -34.0992 | -4.3186
2.3318 | -34.0996 | -4.3276
2.3379 | -34.1001 | -4.339
2.344 | -34.1004 | -4.3499
2.3501 | -34.1009 | -4.3606
2.3562 | -34.1012 | -4.369
2.3623 | -34.1015 | -4.3776
2.3685 | -34.1017 | -4.3866
2.3746 | -34.1019 | -4.3971
2.3807 | -34.1024 | -4.406
2.3868 | -34.1026 | -4.415
2.3929 | -34.103 | -4.4244
2.399 | -34.1033 | -4.4323
2.4051 | -34.1038 | -4.4418
2.4112 | -34.1041 | -4.4508
2.4173 | -34.1044 | -4.4597
2.4234 | -34.1044 | -4.4682
2.4295 | -34.1047 | -4.4768
2.4356 | -34.1048 | -4.4861
2.4417 | -34.1048 | -4.4962
2.4478 | -34.1049 | -4.5058
2.4539 | -34.1052 | -4.5136
2.46 | -34.1049 | -4.5234
2.4661 | -34.1049 | -4.5327
2.4722 | -34.1052 | -4.5414
2.4783 | -34.1053 | -4.5497
2.4844 | -34.1053 | -4.5583
2.4905 | -34.1052 | -4.5675
2.4966 | -34.1049 | -4.5764
2.5027 | -34.1048 | -4.5845
2.5089 | -34.1047 | -4.593
2.515 | -34.1045 | -4.6023
2.5211 | -34.1045 | -4.6106
2.5272 | -34.1041 | -4.6197
2.5333 | -34.1039 | -4.6293
2.5394 | -34.1039 | -4.6383
2.5455 | -34.1037 | -4.6478
2.5516 | -34.1036 | -4.6566
2.5577 | -34.1036 | -4.6666
2.5638 | -34.1033 | -4.6771
2.5699 | -34.1031 | -4.6874
2.576 | -34.1028 | -4.698
2.5821 | -34.1026 | -4.708
2.5882 | -34.1024 | -4.7176
2.5943 | -34.1023 | -4.7275
2.6004 | -34.1021 | -4.7375
2.6065 | -34.1019 | -4.746
2.6126 | -34.1018 | -4.7551
2.6187 | -34.1017 | -4.7638
2.6248 | -34.1015 | -4.7729
2.6309 | -34.1014 | -4.7843
2.637 | -34.1014 | -4.7957
2.6431 | -34.1014 | -4.8052
2.6492 | -34.1012 | -4.8153
2.6554 | -34.1012 | -4.8257
2.6615 | -34.101 | -4.8356
2.6676 | -34.101 | -4.8462
2.6737 | -34.1011 | -4.8571
2.6798 | -34.1011 | -4.8671
2.6859 | -34.1011 | -4.8778
2.692 | -34.101 | -4.8876
2.6981 | -34.101 | -4.8989
2.7042 | -34.1011 | -4.9111
2.7103 | -34.101 | -4.9212
2.7164 | -34.1013 | -4.932
2.7225 | -34.1014 | -4.9435
2.7286 | -34.1014 | -4.9547
2.7347 | -34.1015 | -4.9648
2.7408 | -34.1018 | -4.974
2.7469 | -34.1017 | -4.9855
2.753 | -34.102 | -4.9962
2.7591 | -34.1022 | -5.0083
2.7652 | -34.1022 | -5.02
2.7713 | -34.1022 | -5.03
2.7774 | -34.1026 | -5.0415
2.7835 | -34.1027 | -5.0522
2.7896 | -34.1029 | -5.0616
2.7958 | -34.1033 | -5.0719
2.8019 | -34.1038 | -5.0818
2.808 | -34.1041 | -5.0938
2.8141 | -34.1045 | -5.1069
2.8202 | -34.1049 | -5.1191
2.8263 | -34.1054 | -5.1301
2.8324 | -34.1056 | -5.1428
2.8385 | -34.1061 | -5.155
2.8446 | -34.1067 | -5.1667
2.8507 | -34.1071 | -5.179
2.8568 | -34.1076 | -5.1902
2.8629 | -34.108 | -5.2016
2.869 | -34.1085 | -5.213
2.8751 | -34.1089 | -5.2251
2.8812 | -34.1096 | -5.2385
2.8873 | -34.1099 | -5.2499
2.8934 | -34.1103 | -5.2619
2.8995 | -34.1107 | -5.2741
2.9056 | -34.111 | -5.2849
2.9117 | -34.1115 | -5.2966
2.9178 | -34.1122 | -5.3094
2.9239 | -34.1126 | -5.3229
2.93 | -34.1131 | -5.3366
2.9361 | -34.1137 | -5.348
2.9423 | -34.1143 | -5.3601
2.9484 | -34.1148 | -5.3738
2.9545 | -34.1155 | -5.3871
2.9606 | -34.1161 | -5.4007
2.9667 | -34.1166 | -5.4126
2.9728 | -34.117 | -5.425
2.9789 | -34.1175 | -5.4379
2.985 | -34.1179 | -5.4518
2.9911 | -34.1184 | -5.4645
2.9972 | -34.119 | -5.4772
3.0033 | -34.1196 | -5.4905
3.0094 | -34.1202 | -5.5037
3.0155 | -34.1208 | -5.5164
3.0216 | -34.1213 | -5.5299
3.0277 | -34.1218 | -5.5438
3.0338 | -34.1224 | -5.5576
3.0399 | -34.123 | -5.5714
3.046 | -34.1233 | -5.5846
3.0521 | -34.1239 | -5.5972
3.0582 | -34.1242 | -5.6112
3.0643 | -34.1248 | -5.6244
3.0704 | -34.125 | -5.638
3.0765 | -34.1254 | -5.6503
3.0827 | -34.1258 | -5.6636
3.0888 | -34.1261 | -5.6765
3.0949 | -34.1265 | -5.6896
3.101 | -34.1269 | -5.7039
3.1071 | -34.1272 | -5.7182
3.1132 | -34.1277 | -5.7324
3.1193 | -34.1281 | -5.7456
3.1254 | -34.1286 | -5.7579
3.1315 | -34.129 | -5.7724
3.1376 | -34.1292 | -5.7854
3.1437 | -34.1296 | -5.7992
3.1498 | -34.13 | -5.8135
3.1559 | -34.1303 | -5.8269
3.162 | -34.1303 | -5.8412
3.1681 | -34.1304 | -5.8556
3.1742 | -34.1308 | -5.8695
3.1803 | -34.1311 | -5.8828
3.1864 | -34.1315 | -5.8952
3.1925 | -34.1318 | -5.909
3.1986 | -34.1319 | -5.922
3.2047 | -34.132 | -5.9361
3.2108 | -34.1323 | -5.9479
3.2169 | -34.1327 | -5.9618
3.223 | -34.1326 | -5.9754
3.2292 | -34.1326 | -5.9895
3.2353 | -34.1325 | -5.999
3.2414 | -34.1326 | -6.0122
3.2475 | -34.1326 | -6.0233
3.2536 | -34.1329 | -6.0371
3.2597 | -34.1331 | -6.0484
3.2658 | -34.1333 | -6.0592
3.2719 | -34.1334 | -6.0703
3.278 | -34.1334 | -6.0826
3.2841 | -34.1336 | -6.0946
3.2902 | -34.1337 | -6.1072
3.2963 | -34.1338 | -6.119
3.3024 | -34.1337 | -6.1292
3.3085 | -34.1339 | -6.1419
3.3146 | -34.1338 | -6.1529
3.3207 | -34.134 | -6.1651
3.3268 | -34.1341 | -6.1775
3.3329 | -34.1342 | -6.1873
3.339 | -34.1343 | -6.197
3.3451 | -34.1343 | -6.2084
3.3512 | -34.1344 | -6.2179
3.3573 | -34.1346 | -6.2293
3.3634 | -34.1347 | -6.2409
3.3696 | -34.1346 | -6.2528
3.3757 | -34.1346 | -6.2627
3.3818 | -34.1349 | -6.2745
3.3879 | -34.1348 | -6.2849
3.394 | -34.1347 | -6.2954
3.4001 | -34.1348 | -6.3063
3.4062 | -34.1348 | -6.3165
3.4123 | -34.1348 | -6.3263
3.4184 | -34.1348 | -6.3374
3.4245 | -34.1349 | -6.348
3.4306 | -34.1348 | -6.3577
3.4367 | -34.135 | -6.3675
3.4428 | -34.1353 | -6.3766
3.4489 | -34.1352 | -6.3877
3.455 | -34.1351 | -6.3982
3.4611 | -34.1351 | -6.4094
3.4672 | -34.1351 | -6.4208
3.4733 | -34.1351 | -6.431
3.4794 | -34.1351 | -6.4415
3.4855 | -34.1354 | -6.4511
3.4916 | -34.1355 | -6.4608
3.4977 | -34.1354 | -6.4718
3.5038 | -34.1356 | -6.4818
3.5099 | -34.1356 | -6.4914
3.5161 | -34.1357 | -6.5009
3.5222 | -34.1358 | -6.5109
3.5283 | -34.1357 | -6.521
3.5344 | -34.1358 | -6.5318
3.5405 | -34.1359 | -6.5417
3.5466 | -34.136 | -6.5515
3.5527 | -34.1361 | -6.5622
3.5588 | -34.1362 | -6.572
3.5649 | -34.1362 | -6.5822
3.571 | -34.1363 | -6.5924
3.5771 | -34.1364 | -6.6033
3.5832 | -34.1365 | -6.6128
3.5893 | -34.1367 | -6.6227
3.5954 | -34.1368 | -6.6333
3.6015 | -34.1369 | -6.6437
3.6076 | -34.1372 | -6.6533
3.6137 | -34.1372 | -6.6632
3.6198 | -34.1376 | -6.6722
3.6259 | -34.1377 | -6.683
3.632 | -34.1379 | -6.6936
3.6381 | -34.1381 | -6.7035
3.6442 | -34.1383 | -6.714
3.6503 | -34.1385 | -6.723
3.6565 | -34.1388 | -6.7348
3.6626 | -34.1389 | -6.7456
3.6687 | -34.1392 | -6.7562
3.6748 | -34.1395 | -6.7662
3.6809 | -34.1399 | -6.7767
3.687 | -34.1401 | -6.7879
3.6931 | -34.1405 | -6.7979
3.6992 | -34.1407 | -6.8086
3.7053 | -34.1409 | -6.8187
3.7114 | -34.1411 | -6.83
3.7175 | -34.1415 | -6.8399
3.7236 | -34.1417 | -6.8487
3.7297 | -34.1418 | -6.8595
3.7358 | -34.142 | -6.8713
3.7419 | -34.1422 | -6.8829
3.748 | -34.1424 | -6.8937
3.7541 | -34.1425 | -6.9047
3.7602 | -34.1429 | -6.9156
3.7663 | -34.1431 | -6.9258
3.7724 | -34.1434 | -6.9369
3.7785 | -34.1438 | -6.9473
3.7846 | -34.144 | -6.958
3.7907 | -34.1444 | -6.969
3.7969 | -34.1446 | -6.9806
3.803 | -34.145 | -6.9907
3.8091 | -34.1453 | -7.001
3.8152 | -34.1456 | -7.0122
3.8213 | -34.1458 | -7.0239
3.8274 | -34.1462 | -7.0344
3.8335 | -34.1466 | -7.0461
3.8396 | -34.1471 | -7.0577
3.8457 | -34.1474 | -7.068
3.8518 | -34.1479 | -7.0789
3.8579 | -34.1483 | -7.0896
3.864 | -34.1487 | -7.1003
3.8701 | -34.1491 | -7.1114
3.8762 | -34.1494 | -7.1226
3.8823 | -34.1497 | -7.1333
3.8884 | -34.1501 | -7.1437
3.8945 | -34.1506 | -7.1551
3.9006 | -34.151 | -7.1668
3.9067 | -34.1514 | -7.1788
3.9128 | -34.1517 | -7.1901
3.9189 | -34.1522 | -7.2008
3.925 | -34.1524 | -7.2117
3.9311 | -34.1528 | -7.2218
3.9372 | -34.1532 | -7.2334
3.9434 | -34.1536 | -7.2439
3.9495 | -34.154 | -7.2556
3.9556 | -34.1545 | -7.2669
3.9617 | -34.1548 | -7.2778
3.9678 | -34.1552 | -7.2885
3.9739 | -34.1556 | -7.2998
3.98 | -34.156 | -7.3117
3.9861 | -34.1563 | -7.3222
3.9922 | -34.1565 | -7.3339
3.9983 | -34.1569 | -7.344
4.0044 | -34.1572 | -7.3537
4.0105 | -34.1576 | -7.3647
4.0166 | -34.1581 | -7.3758
4.0227 | -34.1585 | -7.3862
4.0288 | -34.1588 | -7.3977
4.0349 | -34.1591 | -7.4076
4.041 | -34.1594 | -7.4186
4.0471 | -34.1597 | -7.4289
4.0532 | -34.1601 | -7.4401
4.0593 | -34.1604 | -7.4517
4.0654 | -34.1608 | -7.4626
4.0715 | -34.1611 | -7.474
4.0776 | -34.1615 | -7.4845
4.0838 | -34.1618 | -7.496
4.0899 | -34.1622 | -7.5073
4.096 | -34.1625 | -7.5177
4.1021 | -34.1628 | -7.5286
4.1082 | -34.1632 | -7.5389
4.1143 | -34.1635 | -7.5488
4.1204 | -34.164 | -7.5596
4.1265 | -34.1643 | -7.57
4.1326 | -34.1647 | -7.5803
4.1387 | -34.1652 | -7.5916
4.1448 | -34.1655 | -7.6026
4.1509 | -34.1657 | -7.6134
4.157 | -34.166 | -7.6247
4.1631 | -34.1663 | -7.636
4.1692 | -34.1667 | -7.6468
4.1753 | -34.1671 | -7.6577
4.1814 | -34.1673 | -7.6675
4.1875 | -34.1676 | -7.679
4.1936 | -34.1679 | -7.6903
4.1997 | -34.1682 | -7.7005
4.2058 | -34.1685 | -7.7121
4.2119 | -34.1687 | -7.7219
4.218 | -34.169 | -7.732
4.2241 | -34.1694 | -7.7416
4.2303 | -34.1698 | -7.7529
4.2364 | -34.1701 | -7.7644
4.2425 | -34.1704 | -7.7749
4.2486 | -34.1706 | -7.7861
4.2547 | -34.1708 | -7.7954
4.2608 | -34.171 | -7.8053
4.2669 | -34.1713 | -7.8157
4.273 | -34.1716 | -7.8266
4.2791 | -34.1718 | -7.8373
4.2852 | -34.172 | -7.8477
4.2913 | -34.1721 | -7.8569
4.2974 | -34.1724 | -7.8679
4.3035 | -34.1726 | -7.8788
4.3096 | -34.1728 | -7.8893
4.3157 | -34.173 | -7.8992
4.3218 | -34.1732 | -7.9092
4.3279 | -34.1735 | -7.9196
4.334 | -34.1737 | -7.9295
4.3401 | -34.1739 | -7.9394
4.3462 | -34.1742 | -7.9503
4.3523 | -34.1745 | -7.9599
4.3584 | -34.1746 | -7.9684
4.3645 | -34.1748 | -7.978
4.3707 | -34.1751 | -7.9873
4.3768 | -34.1752 | -7.9977
4.3829 | -34.1754 | -8.0081
4.389 | -34.1757 | -8.0178
4.3951 | -34.1759 | -8.0273
4.4012 | -34.1761 | -8.0372
4.4073 | -34.1763 | -8.0473
4.4134 | -34.1765 | -8.0575
4.4195 | -34.1768 | -8.0672
4.4256 | -34.1771 | -8.0767
4.4317 | -34.1773 | -8.0863
4.4378 | -34.1775 | -8.0958
4.4439 | -34.1776 | -8.1052
4.45 | -34.1777 | -8.1152
4.4561 | -34.1778 | -8.125
4.4622 | -34.178 | -8.1347
4.4683 | -34.1781 | -8.1439
4.4744 | -34.1783 | -8.1541
4.4805 | -34.1786 | -8.1631
4.4866 | -34.1788 | -8.1724
4.4927 | -34.179 | -8.1827
4.4988 | -34.1792 | -8.1921
4.5049 | -34.1794 | -8.2012
4.511 | -34.1797 | -8.2102
4.5172 | -34.18 | -8.22
4.5233 | -34.1801 | -8.2306
4.5294 | -34.1804 | -8.24
4.5355 | -34.1806 | -8.2494
4.5416 | -34.1809 | -8.2585
4.5477 | -34.181 | -8.2677
4.5538 | -34.1812 | -8.2764
4.5599 | -34.1815 | -8.2859
4.566 | -34.1817 | -8.2957
4.5721 | -34.182 | -8.3049
4.5782 | -34.1823 | -8.3147
4.5843 | -34.1824 | -8.3244
4.5904 | -34.1828 | -8.3337
4.5965 | -34.1831 | -8.3424
4.6026 | -34.1833 | -8.3531
4.6087 | -34.1835 | -8.3621
4.6148 | -34.1838 | -8.3728
4.6209 | -34.1841 | -8.3812
4.627 | -34.1845 | -8.3913
4.6331 | -34.1845 | -8.4002
4.6392 | -34.1848 | -8.4094
4.6453 | -34.185 | -8.4188
4.6514 | -34.1853 | -8.4284
4.6576 | -34.1853 | -8.4389
4.6637 | -34.1856 | -8.4465
4.6698 | -34.1858 | -8.4551
4.6759 | -34.1859 | -8.465
4.682 | -34.1861 | -8.4741
4.6881 | -34.1862 | -8.4829
4.6942 | -34.1865 | -8.4921
4.7003 | -34.1867 | -8.5016
4.7064 | -34.187 | -8.5108
4.7125 | -34.1873 | -8.5203
4.7186 | -34.1874 | -8.5287
4.7247 | -34.1878 | -8.5384
4.7308 | -34.1881 | -8.5477
4.7369 | -34.1883 | -8.5573
4.743 | -34.1885 | -8.5677
4.7491 | -34.1887 | -8.5774
4.7552 | -34.1888 | -8.5866
4.7613 | -34.189 | -8.5963
4.7674 | -34.1893 | -8.6048
4.7735 | -34.1894 | -8.615
4.7796 | -34.1896 | -8.6248
4.7857 | -34.1897 | -8.6346
4.7918 | -34.19 | -8.6432
4.7979 | -34.1903 | -8.6523
4.8041 | -34.1904 | -8.6616
4.8102 | -34.1908 | -8.6706
4.8163 | -34.191 | -8.6801
4.8224 | -34.1911 | -8.6888
4.8285 | -34.1915 | -8.6978
4.8346 | -34.1917 | -8.7074
4.8407 | -34.1919 | -8.7167
4.8468 | -34.1922 | -8.7265
4.8529 | -34.1924 | -8.7364
4.859 | -34.1926 | -8.7459
4.8651 | -34.1928 | -8.7553
4.8712 | -34.193 | -8.7647
4.8773 | -34.1932 | -8.7743
4.8834 | -34.1935 | -8.7841
4.8895 | -34.1938 | -8.7941
4.8956 | -34.194 | -8.8041
4.9017 | -34.1941 | -8.8136
4.9078 | -34.1943 | -8.8235
4.9139 | -34.1946 | -8.8321
4.92 | -34.1948 | -8.8411
4.9261 | -34.1952 | -8.8499
4.9322 | -34.1955 | -8.8591
4.9383 | -34.1959 | -8.8682
4.9445 | -34.1962 | -8.8783
4.9506 | -34.1964 | -8.8872
4.9567 | -34.1966 | -8.8965
4.9628 | -34.1969 | -8.9059
4.9689 | -34.1971 | -8.9158
4.975 | -34.1973 | -8.9252
4.9811 | -34.1976 | -8.935
4.9872 | -34.1978 | -8.9448
4.9933 | -34.1982 | -8.9546
4.9994 | -34.1985 | -8.9645
5.0055 | -34.1989 | -8.9742
5.0116 | -34.1992 | -8.9845
5.0177 | -34.1996 | -8.9947
5.0238 | -34.1999 | -9.0052
5.0299 | -34.2004 | -9.0163
5.036 | -34.2006 | -9.0259
5.0421 | -34.2008 | -9.0356
5.0482 | -34.2011 | -9.046
5.0543 | -34.2015 | -9.0553
5.0604 | -34.2019 | -9.0656
5.0665 | -34.2022 | -9.0753
5.0726 | -34.2025 | -9.0858
5.0787 | -34.2028 | -9.0962
5.0848 | -34.2031 | -9.1063
5.091 | -34.2034 | -9.1162
5.0971 | -34.2038 | -9.1263
5.1032 | -34.2041 | -9.1367
5.1093 | -34.2045 | -9.1467
5.1154 | -34.2047 | -9.1568
5.1215 | -34.2049 | -9.1668
5.1276 | -34.2053 | -9.1767
5.1337 | -34.2058 | -9.1864
5.1398 | -34.206 | -9.1962
5.1459 | -34.2064 | -9.2068
5.152 | -34.2067 | -9.2179
5.1581 | -34.2069 | -9.2286
5.1642 | -34.2072 | -9.2387
5.1703 | -34.2076 | -9.2488
5.1764 | -34.2079 | -9.2588
5.1825 | -34.2082 | -9.2691
5.1886 | -34.2084 | -9.2789
5.1947 | -34.2087 | -9.2898
5.2008 | -34.2092 | -9.2997
5.2069 | -34.2095 | -9.311
5.213 | -34.2099 | -9.3206
5.2191 | -34.2103 | -9.331
5.2252 | -34.2105 | -9.3414
5.2314 | -34.2107 | -9.3522
5.2375 | -34.2112 | -9.3621
5.2436 | -34.2114 | -9.3724
5.2497 | -34.2118 | -9.3823
5.2558 | -34.2122 | -9.3934
5.2619 | -34.2125 | -9.4044
5.268 | -34.2128 | -9.4149
5.2741 | -34.2132 | -9.4262
5.2802 | -34.2135 | -9.4371
5.2863 | -34.2139 | -9.4474
5.2924 | -34.2142 | -9.4587
5.2985 | -34.2146 | -9.4701
5.3046 | -34.2151 | -9.4814
5.3107 | -34.2155 | -9.4924
5.3168 | -34.2159 | -9.5029
5.3229 | -34.2162 | -9.5128
5.329 | -34.2166 | -9.5228
5.3351 | -34.2171 | -9.5331
5.3412 | -34.2175 | -9.5433
5.3473 | -34.2178 | -9.5543
5.3534 | -34.2182 | -9.5652
5.3595 | -34.2185 | -9.576
5.3656 | -34.2189 | -9.5871
5.3717 | -34.2192 | -9.5977
5.3779 | -34.2196 | -9.6088
5.384 | -34.2199 | -9.6196
5.3901 | -34.2203 | -9.6304
5.3962 | -34.2205 | -9.6413
5.4023 | -34.221 | -9.6521
5.4084 | -34.2214 | -9.6626
5.4145 | -34.2217 | -9.6737
5.4206 | -34.2219 | -9.6851
5.4267 | -34.2223 | -9.6965
5.4328 | -34.2227 | -9.7068
5.4389 | -34.2231 | -9.717
5.445 | -34.2234 | -9.7285
5.4511 | -34.2238 | -9.739
5.4572 | -34.2242 | -9.7503
5.4633 | -34.2245 | -9.7607
5.4694 | -34.2248 | -9.7713
5.4755 | -34.2252 | -9.7822
5.4816 | -34.2255 | -9.7929
5.4877 | -34.2259 | -9.8037
5.4938 | -34.2262 | -9.8142
5.4999 | -34.2265 | -9.8249
5.506 | -34.227 | -9.8358
5.5121 | -34.2273 | -9.8469
5.5183 | -34.2276 | -9.8571
5.5244 | -34.2278 | -9.868
5.5305 | -34.2282 | -9.8789
5.5366 | -34.2286 | -9.8894
5.5427 | -34.2289 | -9.9008
5.5488 | -34.2292 | -9.9119
5.5549 | -34.2295 | -9.9227
5.561 | -34.2298 | -9.9341
5.5671 | -34.2303 | -9.9446
5.5732 | -34.2305 | -9.9556
5.5793 | -34.2309 | -9.9672
5.5854 | -34.2314 | -9.9782
5.5915 | -34.2315 | -9.9893
5.5976 | -34.2319 | -10.0004
5.6037 | -34.2322 | -10.0112
5.6098 | -34.2325 | -10.0218
5.6159 | -34.2329 | -10.0328
5.622 | -34.2333 | -10.0433
5.6281 | -34.2336 | -10.0543
5.6342 | -34.2339 | -10.0652
5.6403 | -34.2342 | -10.0756
5.6464 | -34.2346 | -10.0865
5.6525 | -34.2349 | -10.0971
5.6586 | -34.2353 | -10.1071
5.6648 | -34.2356 | -10.1185
5.6709 | -34.236 | -10.1297
5.677 | -34.2363 | -10.1413
5.6831 | -34.2366 | -10.1517
5.6892 | -34.237 | -10.1625
5.6953 | -34.2373 | -10.1733
5.7014 | -34.2376 | -10.1847
5.7075 | -34.2379 | -10.1957
5.7136 | -34.2383 | -10.2063
5.7197 | -34.2386 | -10.217
5.7258 | -34.2388 | -10.2276
5.7319 | -34.2392 | -10.2384
5.738 | -34.2395 | -10.2492
5.7441 | -34.2399 | -10.26
5.7502 | -34.2403 | -10.2703
5.7563 | -34.2406 | -10.2809
5.7624 | -34.2409 | -10.2924
5.7685 | -34.2412 | -10.3031
5.7746 | -34.2414 | -10.3137
5.7807 | -34.2417 | -10.3241
5.7868 | -34.242 | -10.3349
5.7929 | -34.2423 | -10.3455
5.799 | -34.2426 | -10.3562
5.8052 | -34.2429 | -10.367
5.8113 | -34.2432 | -10.3778
5.8174 | -34.2435 | -10.3887
5.8235 | -34.2438 | -10.3992
5.8296 | -34.2441 | -10.4105
5.8357 | -34.2443 | -10.4213
5.8418 | -34.2446 | -10.4322
5.8479 | -34.2449 | -10.4433
5.854 | -34.2452 | -10.4544
5.8601 | -34.2455 | -10.4649
5.8662 | -34.2458 | -10.476
5.8723 | -34.246 | -10.4861
5.8784 | -34.2463 | -10.4968
5.8845 | -34.2467 | -10.5076
5.8906 | -34.2469 | -10.5183
5.8967 | -34.2473 | -10.5286
5.9028 | -34.2476 | -10.5394
5.9089 | -34.2479 | -10.5495
5.915 | -34.2482 | -10.5602
5.9211 | -34.2485 | -10.5709
5.9272 | -34.2488 | -10.5815
5.9333 | -34.2491 | -10.5915
5.9394 | -34.2495 | -10.6019
5.9455 | -34.2498 | -10.6122
5.9517 | -34.2502 | -10.6228
5.9578 | -34.2505 | -10.6338
5.9639 | -34.2508 | -10.644
5.97 | -34.2512 | -10.6548
5.9761 | -34.2515 | -10.6651
5.9822 | -34.2518 | -10.6759
5.9883 | -34.2522 | -10.6863
5.9944 | -34.2525 | -10.697
6.0005 | -34.2529 | -10.7079
6.0066 | -34.2531 | -10.7181
6.0127 | -34.2535 | -10.728
6.0188 | -34.2538 | -10.739
6.0249 | -34.2541 | -10.7495
6.031 | -34.2544 | -10.7598
6.0371 | -34.2548 | -10.7706
6.0432 | -34.2552 | -10.7812
6.0493 | -34.2555 | -10.7915
6.0554 | -34.2559 | -10.802
6.0615 | -34.2562 | -10.8131
6.0676 | -34.2565 | -10.8243
6.0737 | -34.2569 | -10.835
6.0798 | -34.2571 | -10.8454
6.0859 | -34.2575 | -10.856
6.0921 | -34.2579 | -10.8662
6.0982 | -34.2583 | -10.877
6.1043 | -34.2587 | -10.8874
6.1104 | -34.259 | -10.8985
6.1165 | -34.2593 | -10.9082
6.1226 | -34.2597 | -10.9182
6.1287 | -34.2601 | -10.9288
6.1348 | -34.2605 | -10.9394
6.1409 | -34.2608 | -10.9498
6.147 | -34.2612 | -10.9601
6.1531 | -34.2615 | -10.9705
6.1592 | -34.2619 | -10.981
6.1653 | -34.2622 | -10.9913
6.1714 | -34.2625 | -11.0018
6.1775 | -34.2629 | -11.0126
6.1836 | -34.2632 | -11.0229
6.1897 | -34.2635 | -11.0334
6.1958 | -34.2638 | -11.0441
6.2019 | -34.2642 | -11.0549
6.208 | -34.2646 | -11.0658
6.2141 | -34.265 | -11.0765
6.2202 | -34.2653 | -11.0874
6.2263 | -34.2657 | -11.0978
6.2325 | -34.2661 | -11.1077
6.2386 | -34.2665 | -11.1179
6.2447 | -34.2668 | -11.1286
6.2508 | -34.2671 | -11.139
6.2569 | -34.2675 | -11.1498
6.263 | -34.2679 | -11.1599
6.2691 | -34.2682 | -11.1703
6.2752 | -34.2686 | -11.1806
6.2813 | -34.2689 | -11.1911
6.2874 | -34.2693 | -11.2015
6.2935 | -34.2697 | -11.2118
6.2996 | -34.27 | -11.2223
6.3057 | -34.2704 | -11.2327
6.3118 | -34.2707 | -11.243
6.3179 | -34.2711 | -11.2535
6.324 | -34.2715 | -11.2642
6.3301 | -34.2719 | -11.2744
6.3362 | -34.2723 | -11.2849
6.3423 | -34.2727 | -11.2954
6.3484 | -34.2731 | -11.3061
6.3545 | -34.2735 | -11.3159
6.3606 | -34.2739 | -11.326
6.3667 | -34.2743 | -11.3362
6.3728 | -34.2747 | -11.3466
6.379 | -34.275 | -11.357
6.3851 | -34.2754 | -11.3675
6.3912 | -34.2757 | -11.3776
6.3973 | -34.2762 | -11.3878
6.4034 | -34.2765 | -11.3981
6.4095 | -34.2768 | -11.4087
6.4156 | -34.2772 | -11.4188
6.4217 | -34.2776 | -11.4291
6.4278 | -34.2779 | -11.4394
6.4339 | -34.2784 | -11.4498
6.44 | -34.2788 | -11.4598
6.4461 | -34.2791 | -11.4706
6.4522 | -34.2795 | -11.4803
6.4583 | -34.2798 | -11.4908
6.4644 | -34.2801 | -11.5015
6.4705 | -34.2805 | -11.512
6.4766 | -34.2809 | -11.5228
6.4827 | -34.2813 | -11.5336
6.4888 | -34.2816 | -11.5438
6.4949 | -34.282 | -11.5537
6.501 | -34.2823 | -11.5644
6.5071 | -34.2827 | -11.5749
6.5132 | -34.2832 | -11.5851
6.5194 | -34.2835 | -11.5954
6.5255 | -34.2839 | -11.6053
6.5316 | -34.2842 | -11.6158
6.5377 | -34.2846 | -11.6265
6.5438 | -34.285 | -11.6367
6.5499 | -34.2854 | -11.6469
6.556 | -34.2858 | -11.657
6.5621 | -34.2861 | -11.6676
6.5682 | -34.2864 | -11.6777
6.5743 | -34.2868 | -11.6889
6.5804 | -34.2871 | -11.6992
6.5865 | -34.2874 | -11.7098
6.5926 | -34.2879 | -11.7201
6.5987 | -34.2882 | -11.7301
6.6048 | -34.2886 | -11.7411
6.6109 | -34.289 | -11.7514
6.617 | -34.2893 | -11.7627
6.6231 | -34.2896 | -11.7732
6.6292 | -34.2901 | -11.7835
6.6353 | -34.2904 | -11.7949
6.6414 | -34.2909 | -11.8056
6.6475 | -34.2912 | -11.8161
6.6536 | -34.2916 | -11.8261
6.6597 | -34.292 | -11.8362
6.6659 | -34.2924 | -11.8464
6.672 | -34.2928 | -11.8568
6.6781 | -34.2931 | -11.867
6.6842 | -34.2935 | -11.8775
6.6903 | -34.2939 | -11.8878
6.6964 | -34.2943 | -11.8983
6.7025 | -34.2947 | -11.9086
6.7086 | -34.295 | -11.9188
6.7147 | -34.2954 | -11.9293
6.7208 | -34.2957 | -11.9402
6.7269 | -34.2961 | -11.9502
6.733 | -34.2964 | -11.9597
6.7391 | -34.2968 | -11.9703
6.7452 | -34.2973 | -11.9801
6.7513 | -34.2977 | -11.99
6.7574 | -34.2981 | -12.0007
6.7635 | -34.2985 | -12.0105
6.7696 | -34.2989 | -12.0209
6.7757 | -34.2993 | -12.0315
6.7818 | -34.2997 | -12.0419
6.7879 | -34.3001 | -12.0523
6.794 | -34.3005 | -12.0624
6.8001 | -34.3008 | -12.0726
6.8063 | -34.3012 | -12.0831
6.8124 | -34.3015 | -12.0935
6.8185 | -34.3019 | -12.1038
6.8246 | -34.3023 | -12.1141
6.8307 | -34.3026 | -12.1243
6.8368 | -34.303 | -12.135
6.8429 | -34.3034 | -12.1452
6.849 | -34.3038 | -12.1555
6.8551 | -34.3042 | -12.1659
6.8612 | -34.3046 | -12.176
6.8673 | -34.305 | -12.1861
6.8734 | -34.3053 | -12.1972
6.8795 | -34.3057 | -12.2077
6.8856 | -34.3061 | -12.2175
6.8917 | -34.3065 | -12.2273
6.8978 | -34.3068 | -12.2375
6.9039 | -34.3072 | -12.2475
6.91 | -34.3076 | -12.2579
6.9161 | -34.308 | -12.268
6.9222 | -34.3083 | -12.2783
6.9283 | -34.3087 | -12.2882
6.9344 | -34.309 | -12.2983
6.9405 | -34.3094 | -12.3084
6.9466 | -34.3098 | -12.3188
6.9528 | -34.3102 | -12.3291
6.9589 | -34.3106 | -12.3393
6.965 | -34.3109 | -12.3496
6.9711 | -34.3113 | -12.3598
6.9772 | -34.3118 | -12.3698
6.9833 | -34.3121 | -12.3802
6.9894 | -34.3124 | -12.3909
6.9955 | -34.3128 | -12.4011
7.0016 | -34.3132 | -12.4117
7.0077 | -34.3136 | -12.4222
7.0138 | -34.3139 | -12.4328
7.0199 | -34.3143 | -12.4428
7.026 | -34.3146 | -12.4529
7.0321 | -34.3149 | -12.4632
7.0382 | -34.3152 | -12.4731
7.0443 | -34.3157 | -12.4837
7.0504 | -34.316 | -12.4942
7.0565 | -34.3164 | -12.5047
7.0626 | -34.3167 | -12.5147
7.0687 | -34.3171 | -12.5245
7.0748 | -34.3175 | -12.5345
7.0809 | -34.3178 | -12.5448
7.087 | -34.3182 | -12.5552
7.0932 | -34.3185 | -12.5653
7.0993 | -34.3188 | -12.5756
7.1054 | -34.3192 | -12.5855
7.1115 | -34.3195 | -12.5956
7.1176 | -34.3198 | -12.6056
7.1237 | -34.3201 | -12.6158
7.1298 | -34.3204 | -12.6255
7.1359 | -34.3207 | -12.6353
7.142 | -34.321 | -12.6456
7.1481 | -34.3214 | -12.6554
7.1542 | -34.3217 | -12.6654
7.1603 | -34.322 | -12.6755
7.1664 | -34.3224 | -12.6857
7.1725 | -34.3228 | -12.6956
7.1786 | -34.3232 | -12.7058
7.1847 | -34.3235 | -12.7158
7.1908 | -34.3238 | -12.7261
7.1969 | -34.3242 | -12.7361
7.203 | -34.3245 | -12.7462
7.2091 | -34.3248 | -12.7561
7.2152 | -34.3252 | -12.7662
7.2213 | -34.3255 | -12.776
7.2274 | -34.3258 | -12.786
7.2335 | -34.3262 | -12.7961
7.2397 | -34.3265 | -12.8062
7.2458 | -34.3268 | -12.8157
7.2519 | -34.3272 | -12.8258
7.258 | -34.3275 | -12.836
7.2641 | -34.3279 | -12.846
7.2702 | -34.3282 | -12.8564
7.2763 | -34.3286 | -12.8662
7.2824 | -34.329 | -12.8762
7.2885 | -34.3294 | -12.8861
7.2946 | -34.3297 | -12.8963
7.3007 | -34.33 | -12.9064
7.3068 | -34.3303 | -12.9167
7.3129 | -34.3306 | -12.9265
7.319 | -34.331 | -12.9366
7.3251 | -34.3314 | -12.9463
7.3312 | -34.3317 | -12.956
7.3373 | -34.3321 | -12.9657
7.3434 | -34.3325 | -12.9754
7.3495 | -34.3329 | -12.9849
7.3556 | -34.3332 | -12.9953
7.3617 | -34.3335 | -13.0057
7.3678 | -34.3339 | -13.0158
7.3739 | -34.3342 | -13.0254
7.3801 | -34.3345 | -13.0353
7.3862 | -34.3348 | -13.0451
7.3923 | -34.3352 | -13.0551
7.3984 | -34.3355 | -13.0649
7.4045 | -34.3359 | -13.0748
7.4106 | -34.3362 | -13.0846
7.4167 | -34.3365 | -13.0948
7.4228 | -34.3369 | -13.1046
7.4289 | -34.3372 | -13.1143
7.435 | -34.3377 | -13.1239
7.4411 | -34.338 | -13.1335
7.4472 | -34.3383 | -13.1436
7.4533 | -34.3387 | -13.1535
7.4594 | -34.339 | -13.1638
7.4655 | -34.3394 | -13.1736
7.4716 | -34.3397 | -13.1837
7.4777 | -34.3401 | -13.1931
7.4838 | -34.3404 | -13.2032
7.4899 | -34.3408 | -13.2131
7.496 | -34.3412 | -13.223
7.5021 | -34.3415 | -13.2327
7.5082 | -34.3418 | -13.2427
7.5143 | -34.3422 | -13.2525
7.5204 | -34.3425 | -13.2621
7.5266 | -34.3428 | -13.272
7.5327 | -34.3431 | -13.2817
7.5388 | -34.3434 | -13.2916
7.5449 | -34.3437 | -13.3017
7.551 | -34.3441 | -13.3118
7.5571 | -34.3444 | -13.3215
7.5632 | -34.3448 | -13.3313
7.5693 | -34.3451 | -13.3413
7.5754 | -34.3455 | -13.3512
7.5815 | -34.3459 | -13.3606
7.5876 | -34.3461 | -13.3705
7.5937 | -34.3465 | -13.3803
7.5998 | -34.3468 | -13.3902
7.6059 | -34.3471 | -13.4
7.612 | -34.3475 | -13.4098
7.6181 | -34.3478 | -13.4196
7.6242 | -34.3481 | -13.4295
7.6303 | -34.3485 | -13.4392
7.6364 | -34.3488 | -13.4493
7.6425 | -34.3492 | -13.4595
7.6486 | -34.3495 | -13.4693
7.6547 | -34.3499 | -13.4793
7.6608 | -34.3502 | -13.4891
7.667 | -34.3506 | -13.4984
7.6731 | -34.3508 | -13.5084
7.6792 | -34.3511 | -13.5184
7.6853 | -34.3514 | -13.5282
7.6914 | -34.3517 | -13.5382
7.6975 | -34.352 | -13.5482
7.7036 | -34.3524 | -13.5582
7.7097 | -34.3527 | -13.568
7.7158 | -34.353 | -13.5776
7.7219 | -34.3533 | -13.5874
7.728 | -34.3536 | -13.5974
7.7341 | -34.3539 | -13.6073
7.7402 | -34.3543 | -13.617
7.7463 | -34.3546 | -13.6269
7.7524 | -34.3549 | -13.6371
7.7585 | -34.3552 | -13.647
7.7646 | -34.3555 | -13.6566
7.7707 | -34.3558 | -13.6663
7.7768 | -34.356 | -13.676
7.7829 | -34.3563 | -13.6859
7.789 | -34.3566 | -13.6955
7.7951 | -34.3569 | -13.7057
7.8012 | -34.3573 | -13.716
7.8073 | -34.3576 | -13.7257
7.8135 | -34.3579 | -13.7357
7.8196 | -34.3582 | -13.7455
7.8257 | -34.3586 | -13.7557
7.8318 | -34.3588 | -13.766
7.8379 | -34.3592 | -13.7754
7.844 | -34.3594 | -13.7848
7.8501 | -34.3596 | -13.7949
7.8562 | -34.3599 | -13.8054
7.8623 | -34.3603 | -13.8153
7.8684 | -34.3605 | -13.8249
7.8745 | -34.3609 | -13.8346
7.8806 | -34.3611 | -13.844
7.8867 | -34.3615 | -13.8537
7.8928 | -34.3618 | -13.8633
7.8989 | -34.3622 | -13.8727
7.905 | -34.3625 | -13.883
7.9111 | -34.3628 | -13.8928
7.9172 | -34.3631 | -13.9024
7.9233 | -34.3634 | -13.9124
7.9294 | -34.3637 | -13.9223
7.9355 | -34.364 | -13.9317
7.9416 | -34.3643 | -13.9414
7.9477 | -34.3646 | -13.9508
7.9539 | -34.3649 | -13.9602
7.96 | -34.3652 | -13.9703
7.9661 | -34.3655 | -13.9802
7.9722 | -34.3659 | -13.99
7.9783 | -34.3662 | -13.9998
7.9844 | -34.3665 | -14.0101
7.9905 | -34.3668 | -14.0198
7.9966 | -34.3671 | -14.0296
8.0027 | -34.3674 | -14.0397
8.0088 | -34.3677 | -14.0492
8.0149 | -34.368 | -14.0587
8.021 | -34.3683 | -14.0686
8.0271 | -34.3686 | -14.0788
8.0332 | -34.3688 | -14.0886
8.0393 | -34.3692 | -14.0986
8.0454 | -34.3695 | -14.1087
8.0515 | -34.3698 | -14.1184
8.0576 | -34.37 | -14.1285
8.0637 | -34.3703 | -14.1378
8.0698 | -34.3706 | -14.1472
8.0759 | -34.3709 | -14.1567
8.082 | -34.3713 | -14.1667
8.0881 | -34.3715 | -14.1764
8.0942 | -34.3718 | -14.1866
8.1004 | -34.3721 | -14.1968
8.1065 | -34.3725 | -14.2065
8.1126 | -34.3728 | -14.2163
8.1187 | -34.3731 | -14.2262
8.1248 | -34.3734 | -14.2362
8.1309 | -34.3737 | -14.246
8.137 | -34.374 | -14.2557
8.1431 | -34.3743 | -14.2657
8.1492 | -34.3746 | -14.2755
8.1553 | -34.375 | -14.2854
8.1614 | -34.3753 | -14.2955
8.1675 | -34.3755 | -14.3051
8.1736 | -34.3758 | -14.3152
8.1797 | -34.3761 | -14.3252
8.1858 | -34.3764 | -14.335
8.1919 | -34.3767 | -14.3447
8.198 | -34.3769 | -14.3541
8.2041 | -34.3773 | -14.3643
8.2102 | -34.3777 | -14.3742
8.2163 | -34.378 | -14.3839
8.2224 | -34.3783 | -14.3939
8.2285 | -34.3786 | -14.4035
8.2346 | -34.3789 | -14.4132
8.2408 | -34.3792 | -14.4231
8.2469 | -34.3795 | -14.4328
8.253 | -34.3798 | -14.4426
8.2591 | -34.3801 | -14.4521
8.2652 | -34.3805 | -14.4618
8.2713 | -34.3808 | -14.4716
8.2774 | -34.3811 | -14.4815
8.2835 | -34.3815 | -14.4918
8.2896 | -34.3818 | -14.5011
8.2957 | -34.3821 | -14.5109
8.3018 | -34.3824 | -14.5206
8.3079 | -34.3827 | -14.5304
8.314 | -34.383 | -14.5401
8.3201 | -34.3834 | -14.5498
8.3262 | -34.3836 | -14.5595
8.3323 | -34.384 | -14.5693
8.3384 | -34.3843 | -14.5795
8.3445 | -34.3846 | -14.5889
8.3506 | -34.3849 | -14.5989
8.3567 | -34.3853 | -14.609
8.3628 | -34.3855 | -14.6187
8.3689 | -34.3858 | -14.6286
8.375 | -34.3861 | -14.6384
8.3812 | -34.3865 | -14.6482
8.3873 | -34.3868 | -14.6581
8.3934 | -34.3871 | -14.6678
8.3995 | -34.3874 | -14.6777
8.4056 | -34.3877 | -14.6876
8.4117 | -34.388 | -14.6977
8.4178 | -34.3883 | -14.7071
8.4239 | -34.3886 | -14.7166
8.43 | -34.389 | -14.7264
8.4361 | -34.3893 | -14.7366
8.4422 | -34.3896 | -14.7464
8.4483 | -34.3898 | -14.7562
8.4544 | -34.3902 | -14.7662
8.4605 | -34.3904 | -14.7759
8.4666 | -34.3907 | -14.7858
8.4727 | -34.391 | -14.7956
8.4788 | -34.3913 | -14.8052
8.4849 | -34.3916 | -14.8154
8.491 | -34.3919 | -14.825
8.4971 | -34.3922 | -14.8349
8.5032 | -34.3926 | -14.8447
8.5093 | -34.3929 | -14.8544
8.5154 | -34.3932 | -14.864
8.5215 | -34.3935 | -14.8739
8.5277 | -34.3938 | -14.8838
8.5338 | -34.3941 | -14.8932
8.5399 | -34.3944 | -14.9033
8.546 | -34.3948 | -14.9129
8.5521 | -34.3951 | -14.9228
8.5582 | -34.3954 | -14.9328
8.5643 | -34.3957 | -14.9424
8.5704 | -34.396 | -14.9518
8.5765 | -34.3963 | -14.9616
8.5826 | -34.3967 | -14.9714
8.5887 | -34.397 | -14.9811
8.5948 | -34.3973 | -14.9906
8.6009 | -34.3976 | -15.0003
8.607 | -34.3979 | -15.0103
8.6131 | -34.3982 | -15.0203
8.6192 | -34.3985 | -15.0298
8.6253 | -34.3988 | -15.0397
8.6314 | -34.3992 | -15.0495
8.6375 | -34.3995 | -15.0594
8.6436 | -34.3999 | -15.0693
8.6497 | -34.4002 | -15.0792
8.6558 | -34.4005 | -15.0888
8.6619 | -34.4008 | -15.0988
8.6681 | -34.401 | -15.1085
8.6742 | -34.4014 | -15.1186
8.6803 | -34.4017 | -15.1286
8.6864 | -34.4021 | -15.1387
8.6925 | -34.4024 | -15.1484
8.6986 | -34.4027 | -15.1585
8.7047 | -34.4031 | -15.1684
8.7108 | -34.4034 | -15.1783
8.7169 | -34.4037 | -15.1879
8.723 | -34.404 | -15.1977
8.7291 | -34.4043 | -15.2075
8.7352 | -34.4047 | -15.2176
8.7413 | -34.405 | -15.2274
8.7474 | -34.4053 | -15.2369
8.7535 | -34.4057 | -15.2468
8.7596 | -34.406 | -15.2571
8.7657 | -34.4063 | -15.2669
8.7718 | -34.4066 | -15.2765
8.7779 | -34.407 | -15.2868
8.784 | -34.4072 | -15.2965
8.7901 | -34.4075 | -15.3068
8.7962 | -34.4079 | -15.3167
8.8023 | -34.4083 | -15.3262
8.8084 | -34.4086 | -15.3362
8.8146 | -34.4089 | -15.3459
8.8207 | -34.4092 | -15.3556
8.8268 | -34.4096 | -15.3654
8.8329 | -34.4099 | -15.3754
8.839 | -34.4103 | -15.3854
8.8451 | -34.4106 | -15.3949
8.8512 | -34.4109 | -15.4048
8.8573 | -34.4113 | -15.4145
8.8634 | -34.4117 | -15.4243
8.8695 | -34.412 | -15.4343
8.8756 | -34.4123 | -15.4439
8.8817 | -34.4127 | -15.454
8.8878 | -34.4129 | -15.4641
8.8939 | -34.4132 | -15.4736
8.9 | -34.4136 | -15.4832
8.9061 | -34.4139 | -15.493
8.9122 | -34.4143 | -15.5027
8.9183 | -34.4146 | -15.5125
8.9244 | -34.4149 | -15.5221
8.9305 | -34.4153 | -15.5319
8.9366 | -34.4157 | -15.5415
8.9427 | -34.416 | -15.5515
8.9488 | -34.4164 | -15.5612
8.955 | -34.4167 | -15.571
8.9611 | -34.417 | -15.581
8.9672 | -34.4174 | -15.5913
8.9733 | -34.4178 | -15.601
8.9794 | -34.4181 | -15.611
8.9855 | -34.4185 | -15.6207
8.9916 | -34.4188 | -15.6304
8.9977 | -34.4191 | -15.6402
9.0038 | -34.4195 | -15.6502
9.0099 | -34.4199 | -15.6602
9.016 | -34.4202 | -15.67
9.0221 | -34.4206 | -15.6801
9.0282 | -34.4209 | -15.69
9.0343 | -34.4213 | -15.7001
9.0404 | -34.4217 | -15.7099
9.0465 | -34.4221 | -15.7201
9.0526 | -34.4224 | -15.7299
9.0587 | -34.4228 | -15.7397
9.0648 | -34.4232 | -15.7495
9.0709 | -34.4236 | -15.7591
9.077 | -34.4239 | -15.7696
9.0831 | -34.4242 | -15.7794
9.0892 | -34.4245 | -15.7893
9.0953 | -34.4249 | -15.7992
9.1015 | -34.4252 | -15.8091
9.1076 | -34.4256 | -15.819
9.1137 | -34.426 | -15.8288
9.1198 | -34.4264 | -15.8383
9.1259 | -34.4267 | -15.8482
9.132 | -34.4271 | -15.8584
9.1381 | -34.4275 | -15.8681
9.1442 | -34.428 | -15.8776
9.1503 | -34.4282 | -15.8876
9.1564 | -34.4286 | -15.8974
9.1625 | -34.4289 | -15.9074
9.1686 | -34.4294 | -15.9172
9.1747 | -34.4297 | -15.9272
9.1808 | -34.4301 | -15.9373
9.1869 | -34.4304 | -15.9475
9.193 | -34.4308 | -15.9575
9.1991 | -34.4311 | -15.9671
9.2052 | -34.4315 | -15.9769
9.2113 | -34.4319 | -15.987
9.2174 | -34.4323 | -15.9976
9.2235 | -34.4327 | -16.0071
9.2296 | -34.4331 | -16.0172
9.2357 | -34.4335 | -16.0271
9.2419 | -34.4338 | -16.0369
9.248 | -34.4342 | -16.0471
9.2541 | -34.4346 | -16.0571
9.2602 | -34.4349 | -16.0671
9.2663 | -34.4353 | -16.0768
9.2724 | -34.4357 | -16.0868
9.2785 | -34.4361 | -16.0967
9.2846 | -34.4365 | -16.1066
9.2907 | -34.4368 | -16.1164
9.2968 | -34.4371 | -16.1265
9.3029 | -34.4375 | -16.1365
9.309 | -34.4379 | -16.1464
9.3151 | -34.4383 | -16.1562
9.3212 | -34.4387 | -16.1662
9.3273 | -34.4391 | -16.176
9.3334 | -34.4395 | -16.1856
9.3395 | -34.44 | -16.1954
9.3456 | -34.4404 | -16.2051
9.3517 | -34.4408 | -16.2146
9.3578 | -34.4412 | -16.2244
9.3639 | -34.4416 | -16.2345
9.37 | -34.442 | -16.2445
9.3761 | -34.4423 | -16.2542
9.3822 | -34.4427 | -16.264
9.3884 | -34.4431 | -16.2737
9.3945 | -34.4435 | -16.2837
9.4006 | -34.4439 | -16.2937
9.4067 | -34.4443 | -16.3037
9.4128 | -34.4447 | -16.3135
9.4189 | -34.4451 | -16.3236
9.425 | -34.4455 | -16.3333
9.4311 | -34.4459 | -16.3432
9.4372 | -34.4462 | -16.3531
9.4433 | -34.4466 | -16.3632
9.4494 | -34.447 | -16.373
9.4555 | -34.4474 | -16.3828
9.4616 | -34.4478 | -16.3926
9.4677 | -34.4483 | -16.4023
9.4738 | -34.4486 | -16.4126
9.4799 | -34.449 | -16.4224
9.486 | -34.4493 | -16.4327
9.4921 | -34.4497 | -16.443
9.4982 | -34.4501 | -16.4525
9.5043 | -34.4505 | -16.462
9.5104 | -34.4509 | -16.4717
9.5165 | -34.4513 | -16.4816
9.5226 | -34.4517 | -16.4916
9.5288 | -34.4521 | -16.5014
9.5349 | -34.4525 | -16.5113
9.541 | -34.4529 | -16.5211
9.5471 | -34.4533 | -16.531
9.5532 | -34.4537 | -16.5409
9.5593 | -34.4542 | -16.5508
9.5654 | -34.4546 | -16.5607
9.5715 | -34.455 | -16.5708
9.5776 | -34.4553 | -16.581
9.5837 | -34.4557 | -16.5906
9.5898 | -34.4561 | -16.6007
9.5959 | -34.4564 | -16.6105
9.602 | -34.4568 | -16.6203
9.6081 | -34.4572 | -16.6304
9.6142 | -34.4576 | -16.6401
9.6203 | -34.458 | -16.65
9.6264 | -34.4583 | -16.6594
9.6325 | -34.4586 | -16.6695
9.6386 | -34.4591 | -16.6793
9.6447 | -34.4594 | -16.6891
9.6508 | -34.4598 | -16.6991
9.6569 | -34.4602 | -16.7089
9.663 | -34.4607 | -16.7188
9.6691 | -34.4611 | -16.7287
9.6753 | -34.4615 | -16.7386
9.6814 | -34.4619 | -16.7481
9.6875 | -34.4624 | -16.7578
9.6936 | -34.4628 | -16.7675
9.6997 | -34.4632 | -16.7776
9.7058 | -34.4636 | -16.7875
9.7119 | -34.464 | -16.7972
9.718 | -34.4644 | -16.8074
9.7241 | -34.4647 | -16.8173
9.7302 | -34.4651 | -16.8272
9.7363 | -34.4655 | -16.8369
9.7424 | -34.4659 | -16.8467
9.7485 | -34.4663 | -16.8564
9.7546 | -34.4667 | -16.8659
9.7607 | -34.4671 | -16.876
9.7668 | -34.4675 | -16.8863
9.7729 | -34.4679 | -16.8959
9.779 | -34.4684 | -16.9058
9.7851 | -34.4688 | -16.9157
9.7912 | -34.4692 | -16.926
9.7973 | -34.4695 | -16.9362
9.8034 | -34.47 | -16.9462
9.8095 | -34.4703 | -16.956
9.8157 | -34.4707 | -16.9659
9.8218 | -34.4711 | -16.9754
9.8279 | -34.4716 | -16.9852
9.834 | -34.472 | -16.995
9.8401 | -34.4724 | -17.0047
9.8462 | -34.4728 | -17.0143
9.8523 | -34.4732 | -17.024
9.8584 | -34.4736 | -17.034
9.8645 | -34.474 | -17.0439
9.8706 | -34.4744 | -17.0535
9.8767 | -34.4748 | -17.0629
9.8828 | -34.4752 | -17.0723
9.8889 | -34.4756 | -17.0824
9.895 | -34.476 | -17.0919
9.9011 | -34.4764 | -17.1013
9.9072 | -34.4768 | -17.1108
9.9133 | -34.4772 | -17.1205
9.9194 | -34.4776 | -17.1301
9.9255 | -34.478 | -17.1399
9.9316 | -34.4785 | -17.1494
9.9377 | -34.4789 | -17.159
9.9438 | -34.4793 | -17.1687
9.9499 | -34.4797 | -17.1786
9.956 | -34.4801 | -17.1881
9.9622 | -34.4805 | -17.1977
9.9683 | -34.4809 | -17.2075
9.9744 | -34.4813 | -17.2173
9.9805 | -34.4817 | -17.2271
9.9866 | -34.4821 | -17.2365
9.9927 | -34.4826 | -17.2463
9.9988 | -34.483 | -17.2561
10.0049 | -34.4834 | -17.2657
10.011 | -34.4839 | -17.2753
10.0171 | -34.4842 | -17.285
10.0232 | -34.4847 | -17.2945
10.0293 | -34.4852 | -17.304
10.0354 | -34.4856 | -17.3133
10.0415 | -34.486 | -17.3229
10.0476 | -34.4864 | -17.3329
10.0537 | -34.4868 | -17.343
10.0598 | -34.4872 | -17.3528
10.0659 | -34.4876 | -17.3625
10.072 | -34.488 | -17.3718
10.0781 | -34.4885 | -17.3814
10.0842 | -34.4889 | -17.3912
10.0903 | -34.4893 | -17.4011
10.0964 | -34.4898 | -17.4105
10.1026 | -34.4902 | -17.42
10.1087 | -34.4906 | -17.4297
10.1148 | -34.4911 | -17.4395
10.1209 | -34.4915 | -17.449
10.127 | -34.4919 | -17.4584
10.1331 | -34.4923 | -17.4679
10.1392 | -34.4928 | -17.4773
10.1453 | -34.4932 | -17.487
10.1514 | -34.4936 | -17.4963
10.1575 | -34.494 | -17.506
10.1636 | -34.4945 | -17.5149
10.1697 | -34.4948 | -17.5246
10.1758 | -34.4953 | -17.5341
10.1819 | -34.4957 | -17.5439
10.188 | -34.4961 | -17.5532
10.1941 | -34.4965 | -17.5629
10.2002 | -34.4968 | -17.5719
10.2063 | -34.4973 | -17.5816
10.2124 | -34.4977 | -17.5914
10.2185 | -34.4981 | -17.6009
10.2246 | -34.4985 | -17.6104
10.2307 | -34.4989 | -17.6201
10.2368 | -34.4994 | -17.6288
10.2429 | -34.4998 | -17.6381
10.2491 | -34.5002 | -17.6475
10.2552 | -34.5006 | -17.6574
10.2613 | -34.5011 | -17.6672
10.2674 | -34.5014 | -17.6767
10.2735 | -34.5018 | -17.6865
10.2796 | -34.5022 | -17.6962
10.2857 | -34.5026 | -17.7058
10.2918 | -34.503 | -17.7155
10.2979 | -34.5035 | -17.7253
10.304 | -34.5039 | -17.7348
10.3101 | -34.5043 | -17.7442
10.3162 | -34.5047 | -17.7537
10.3223 | -34.5051 | -17.7632
10.3284 | -34.5055 | -17.7726
10.3345 | -34.506 | -17.7818
10.3406 | -34.5063 | -17.7911
10.3467 | -34.5068 | -17.8006
10.3528 | -34.5071 | -17.8098
10.3589 | -34.5075 | -17.8195
10.365 | -34.5079 | -17.8292
10.3711 | -34.5083 | -17.8389
10.3772 | -34.5086 | -17.8484
10.3833 | -34.509 | -17.8578
10.3895 | -34.5094 | -17.8672
10.3956 | -34.5099 | -17.8765
10.4017 | -34.5102 | -17.8858
10.4078 | -34.5106 | -17.8951
10.4139 | -34.5111 | -17.9045
10.42 | -34.5115 | -17.9141
10.4261 | -34.5119 | -17.9234
10.4322 | -34.5124 | -17.9325
10.4383 | -34.5127 | -17.9419
10.4444 | -34.5132 | -17.9509
10.4505 | -34.5136 | -17.9602
10.4566 | -34.514 | -17.9696
10.4627 | -34.5144 | -17.979
10.4688 | -34.5149 | -17.9879
10.4749 | -34.5153 | -17.9972
10.481 | -34.5157 | -18.0067
10.4871 | -34.5161 | -18.0162
10.4932 | -34.5166 | -18.0257
10.4993 | -34.517 | -18.0351
10.5054 | -34.5174 | -18.0446
10.5115 | -34.5179 | -18.054
10.5176 | -34.5183 | -18.0631
10.5237 | -34.5187 | -18.0724
10.5298 | -34.5191 | -18.0814
10.536 | -34.5195 | -18.0908
10.5421 | -34.5199 | -18.0997
10.5482 | -34.5203 | -18.1093
10.5543 | -34.5208 | -18.1187
10.5604 | -34.5212 | -18.1277
10.5665 | -34.5216 | -18.1369
10.5726 | -34.522 | -18.1462
10.5787 | -34.5224 | -18.1557
10.5848 | -34.5229 | -18.1643
10.5909 | -34.5233 | -18.1738
10.597 | -34.5237 | -18.183
10.6031 | -34.5242 | -18.1924
10.6092 | -34.5246 | -18.2021
10.6153 | -34.525 | -18.2118
10.6214 | -34.5255 | -18.2205
10.6275 | -34.5259 | -18.2295
10.6336 | -34.5263 | -18.239
10.6397 | -34.5267 | -18.2483
10.6458 | -34.5271 | -18.2574
10.6519 | -34.5275 | -18.2666
10.658 | -34.5279 | -18.2758
10.6641 | -34.5283 | -18.2852
10.6702 | -34.5288 | -18.2949
10.6764 | -34.5292 | -18.3042
10.6825 | -34.5297 | -18.3135
10.6886 | -34.5301 | -18.3229
10.6947 | -34.5305 | -18.3318
10.7008 | -34.531 | -18.3409
10.7069 | -34.5314 | -18.3501
10.713 | -34.5318 | -18.3592
10.7191 | -34.5323 | -18.3689
10.7252 | -34.5327 | -18.3775
10.7313 | -34.5332 | -18.3868
10.7374 | -34.5335 | -18.3959
10.7435 | -34.5339 | -18.4054
10.7496 | -34.5344 | -18.4145
10.7557 | -34.5348 | -18.4236
10.7618 | -34.5352 | -18.4327
10.7679 | -34.5357 | -18.4419
10.774 | -34.536 | -18.4511
10.7801 | -34.5364 | -18.4608
10.7862 | -34.5369 | -18.4699
10.7923 | -34.5373 | -18.4792
10.7984 | -34.5377 | -18.4883
10.8045 | -34.5381 | -18.4976
10.8106 | -34.5385 | -18.5068
10.8168 | -34.5389 | -18.516
10.8229 | -34.5393 | -18.5253
10.829 | -34.5398 | -18.5345
10.8351 | -34.5402 | -18.5434
10.8412 | -34.5406 | -18.5525
10.8473 | -34.541 | -18.5615
10.8534 | -34.5414 | -18.5707
10.8595 | -34.5418 | -18.5797
10.8656 | -34.5423 | -18.5887
10.8717 | -34.5427 | -18.5975
10.8778 | -34.5432 | -18.6064
10.8839 | -34.5436 | -18.6153
10.89 | -34.544 | -18.6242
10.8961 | -34.5444 | -18.6335
10.9022 | -34.5448 | -18.6422
10.9083 | -34.5453 | -18.6516
10.9144 | -34.5457 | -18.6606
10.9205 | -34.5462 | -18.6697
10.9266 | -34.5466 | -18.6791
10.9327 | -34.5471 | -18.6881
10.9388 | -34.5475 | -18.6973
10.9449 | -34.5479 | -18.7065
10.951 | -34.5484 | -18.7156
10.9571 | -34.5488 | -18.7249
10.9633 | -34.5492 | -18.7343
10.9694 | -34.5496 | -18.7437
10.9755 | -34.5501 | -18.7528
10.9816 | -34.5505 | -18.762
10.9877 | -34.5509 | -18.7711
10.9938 | -34.5513 | -18.7804
10.9999 | -34.5518 | -18.7895
11.006 | -34.5523 | -18.7984
11.0121 | -34.5526 | -18.8079
11.0182 | -34.5531 | -18.8167
11.0243 | -34.5535 | -18.8257
11.0304 | -34.5539 | -18.8346
11.0365 | -34.5543 | -18.8437
11.0426 | -34.5548 | -18.8526
11.0487 | -34.5552 | -18.8616
11.0548 | -34.5557 | -18.8704
11.0609 | -34.5561 | -18.8797
11.067 | -34.5565 | -18.889
11.0731 | -34.5569 | -18.8981
11.0792 | -34.5573 | -18.9072
11.0853 | -34.5577 | -18.9158
11.0914 | -34.5582 | -18.9249
11.0975 | -34.5586 | -18.9343
11.1037 | -34.559 | -18.943
11.1098 | -34.5594 | -18.9524
11.1159 | -34.5598 | -18.9617
11.122 | -34.5603 | -18.9709
11.1281 | -34.5607 | -18.9797
11.1342 | -34.5611 | -18.9885
11.1403 | -34.5615 | -18.9981
11.1464 | -34.562 | -19.007
11.1525 | -34.5624 | -19.016
11.1586 | -34.5628 | -19.0257
11.1647 | -34.5633 | -19.0344
11.1708 | -34.5637 | -19.0432
11.1769 | -34.5641 | -19.0518
11.183 | -34.5645 | -19.0606
11.1891 | -34.565 | -19.0697
11.1952 | -34.5654 | -19.0792
11.2013 | -34.5659 | -19.0881
11.2074 | -34.5663 | -19.0969
11.2135 | -34.5666 | -19.1062
11.2196 | -34.567 | -19.115
11.2257 | -34.5674 | -19.1238
11.2318 | -34.5678 | -19.133
11.2379 | -34.5682 | -19.1418
11.244 | -34.5686 | -19.1508
11.2502 | -34.5691 | -19.1598
11.2563 | -34.5696 | -19.1689
11.2624 | -34.5699 | -19.178
11.2685 | -34.5704 | -19.1867
11.2746 | -34.5708 | -19.1954
11.2807 | -34.5712 | -19.205
11.2868 | -34.5716 | -19.2142
11.2929 | -34.572 | -19.2228
11.299 | -34.5725 | -19.2321
11.3051 | -34.5728 | -19.2412
11.3112 | -34.5732 | -19.2502
11.3173 | -34.5737 | -19.2593
11.3234 | -34.5741 | -19.2683
11.3295 | -34.5746 | -19.277
11.3356 | -34.575 | -19.286
11.3417 | -34.5754 | -19.295
11.3478 | -34.5758 | -19.304
11.3539 | -34.5762 | -19.3129
11.36 | -34.5766 | -19.3218
11.3661 | -34.5771 | -19.3306
11.3722 | -34.5775 | -19.3392
11.3783 | -34.5779 | -19.348
11.3844 | -34.5783 | -19.3569
11.3906 | -34.5787 | -19.3662
11.3967 | -34.5791 | -19.3749
11.4028 | -34.5796 | -19.3842
11.4089 | -34.58 | -19.3929
11.415 | -34.5804 | -19.4018
11.4211 | -34.5808 | -19.4104
11.4272 | -34.5812 | -19.4194
11.4333 | -34.5816 | -19.4284
11.4394 | -34.582 | -19.4377
11.4455 | -34.5824 | -19.4465
11.4516 | -34.5829 | -19.4554
11.4577 | -34.5833 | -19.4645
11.4638 | -34.5837 | -19.4731
11.4699 | -34.5841 | -19.4825
11.476 | -34.5845 | -19.492
11.4821 | -34.585 | -19.5006
11.4882 | -34.5854 | -19.5099
11.4943 | -34.5858 | -19.5187
11.5004 | -34.5863 | -19.5277
11.5065 | -34.5868 | -19.5365
11.5126 | -34.5872 | -19.5451
11.5187 | -34.5876 | -19.5541
11.5248 | -34.5881 | -19.5631
11.5309 | -34.5885 | -19.5719
11.5371 | -34.5889 | -19.5808
11.5432 | -34.5893 | -19.5901
11.5493 | -34.5897 | -19.5991
11.5554 | -34.5901 | -19.6079
11.5615 | -34.5905 | -19.6172
11.5676 | -34.591 | -19.6263
11.5737 | -34.5914 | -19.6349
11.5798 | -34.5918 | -19.6435
11.5859 | -34.5923 | -19.6525
11.592 | -34.5927 | -19.6614
11.5981 | -34.5931 | -19.6701
11.6042 | -34.5935 | -19.6791
11.6103 | -34.5939 | -19.6878
11.6164 | -34.5944 | -19.6965
11.6225 | -34.5948 | -19.7055
11.6286 | -34.5952 | -19.7144
11.6347 | -34.5956 | -19.7236
11.6408 | -34.5961 | -19.7326
11.6469 | -34.5966 | -19.7418
11.653 | -34.597 | -19.7505
11.6591 | -34.5975 | -19.7593
11.6652 | -34.5979 | -19.7683
11.6713 | -34.5983 | -19.7772
11.6775 | -34.5988 | -19.7859
11.6836 | -34.5992 | -19.7949
11.6897 | -34.5997 | -19.8036
11.6958 | -34.6001 | -19.8129
11.7019 | -34.6006 | -19.8218
11.708 | -34.601 | -19.8308
11.7141 | -34.6015 | -19.8397
11.7202 | -34.6019 | -19.8483
11.7263 | -34.6023 | -19.8572
11.7324 | -34.6027 | -19.8663
11.7385 | -34.6032 | -19.8754
11.7446 | -34.6036 | -19.8845
11.7507 | -34.604 | -19.8932
11.7568 | -34.6044 | -19.9022
11.7629 | -34.6049 | -19.9113
11.769 | -34.6053 | -19.9202
11.7751 | -34.6058 | -19.9288
11.7812 | -34.6062 | -19.9375
11.7873 | -34.6066 | -19.9466
11.7934 | -34.6071 | -19.9555
11.7995 | -34.6075 | -19.9642
11.8056 | -34.6079 | -19.9735
11.8117 | -34.6083 | -19.9824
11.8178 | -34.6088 | -19.9912
11.824 | -34.6093 | -19.9999
11.8301 | -34.6098 | -20.0085
11.8362 | -34.6101 | -20.017
11.8423 | -34.6105 | -20.0259
11.8484 | -34.611 | -20.0347
11.8545 | -34.6114 | -20.0438
11.8606 | -34.6119 | -20.0527
11.8667 | -34.6123 | -20.0617
11.8728 | -34.6127 | -20.0706
11.8789 | -34.6131 | -20.0799
11.885 | -34.6135 | -20.0887
11.8911 | -34.614 | -20.0977
11.8972 | -34.6144 | -20.1063
11.9033 | -34.6149 | -20.1149
11.9094 | -34.6152 | -20.1234
11.9155 | -34.6156 | -20.1323
11.9216 | -34.6161 | -20.141
11.9277 | -34.6166 | -20.1495
11.9338 | -34.617 | -20.1585
11.9399 | -34.6174 | -20.1677
11.946 | -34.6178 | -20.1765
11.9521 | -34.6183 | -20.1851
11.9582 | -34.6187 | -20.1939
11.9644 | -34.6192 | -20.2026
11.9705 | -34.6196 | -20.2116
11.9766 | -34.6201 | -20.2205
11.9827 | -34.6205 | -20.2292
11.9888 | -34.6209 | -20.2379
11.9949 | -34.6214 | -20.2469
12.001 | -34.6219 | -20.2557
12.0071 | -34.6223 | -20.2644
12.0132 | -34.6228 | -20.2729
12.0193 | -34.6232 | -20.2822
12.0254 | -34.6237 | -20.2908
12.0315 | -34.6242 | -20.2998
12.0376 | -34.6246 | -20.3084
12.0437 | -34.625 | -20.3171
12.0498 | -34.6254 | -20.3261
12.0559 | -34.6259 | -20.3351
12.062 | -34.6263 | -20.3441
12.0681 | -34.6268 | -20.3529
12.0742 | -34.6272 | -20.3619
12.0803 | -34.6277 | -20.3706
12.0864 | -34.6281 | -20.3795
12.0925 | -34.6286 | -20.3884
12.0986 | -34.6291 | -20.397
12.1047 | -34.6295 | -20.4057
12.1109 | -34.63 | -20.4148
12.117 | -34.6305 | -20.4233
12.1231 | -34.6309 | -20.4322
12.1292 | -34.6313 | -20.4412
12.1353 | -34.6318 | -20.4498
12.1414 | -34.6323 | -20.4588
12.1475 | -34.6327 | -20.4677
12.1536 | -34.6331 | -20.4765
12.1597 | -34.6336 | -20.4855
12.1658 | -34.634 | -20.4943
12.1719 | -34.6345 | -20.5034
12.178 | -34.635 | -20.5122
12.1841 | -34.6354 | -20.5212
12.1902 | -34.6359 | -20.53
12.1963 | -34.6364 | -20.5389
12.2024 | -34.6369 | -20.5477
12.2085 | -34.6373 | -20.5565
12.2146 | -34.6377 | -20.5652
12.2207 | -34.6382 | -20.574
12.2268 | -34.6386 | -20.5829
12.2329 | -34.6391 | -20.5917
12.239 | -34.6395 | -20.6002
12.2451 | -34.6399 | -20.6088
12.2513 | -34.6404 | -20.6177
12.2574 | -34.6409 | -20.6268
12.2635 | -34.6414 | -20.6352
12.2696 | -34.6418 | -20.6438
12.2757 | -34.6422 | -20.6524
12.2818 | -34.6427 | -20.661
12.2879 | -34.6431 | -20.6698
12.294 | -34.6435 | -20.6783
12.3001 | -34.644 | -20.6873
12.3062 | -34.6445 | -20.6962
12.3123 | -34.6449 | -20.7051
12.3184 | -34.6454 | -20.7136
12.3245 | -34.6459 | -20.7225
12.3306 | -34.6463 | -20.7311
12.3367 | -34.6468 | -20.7397
12.3428 | -34.6473 | -20.7483
12.3489 | -34.6477 | -20.7566
12.355 | -34.6482 | -20.7654
12.3611 | -34.6486 | -20.7742
12.3672 | -34.649 | -20.7828
12.3733 | -34.6495 | -20.7919
12.3794 | -34.6499 | -20.8004
12.3855 | -34.6504 | -20.8089
12.3916 | -34.6508 | -20.8177
12.3978 | -34.6513 | -20.8268
12.4039 | -34.6517 | -20.8355
12.41 | -34.6522 | -20.8444
12.4161 | -34.6527 | -20.8529
12.4222 | -34.6531 | -20.8615
12.4283 | -34.6536 | -20.8703
12.4344 | -34.6541 | -20.8791
12.4405 | -34.6545 | -20.8879
12.4466 | -34.655 | -20.8966
12.4527 | -34.6554 | -20.9054
12.4588 | -34.6559 | -20.9142
12.4649 | -34.6563 | -20.9233
12.471 | -34.6568 | -20.9327
12.4771 | -34.6573 | -20.9409
12.4832 | -34.6577 | -20.9497
12.4893 | -34.6582 | -20.9582
12.4954 | -34.6587 | -20.9669
12.5015 | -34.6592 | -20.9754
12.5076 | -34.6596 | -20.9844
12.5137 | -34.66 | -20.9935
12.5198 | -34.6605 | -21.0021
12.5259 | -34.6609 | -21.0106
12.532 | -34.6614 | -21.0196
12.5382 | -34.6619 | -21.0284
12.5443 | -34.6623 | -21.0373
12.5504 | -34.6628 | -21.0461
12.5565 | -34.6633 | -21.0548
12.5626 | -34.6638 | -21.0639
12.5687 | -34.6642 | -21.0726
12.5748 | -34.6647 | -21.0813
12.5809 | -34.6651 | -21.0898
12.587 | -34.6655 | -21.0991
12.5931 | -34.666 | -21.108
12.5992 | -34.6665 | -21.1169
12.6053 | -34.6669 | -21.1255
12.6114 | -34.6674 | -21.1343
12.6175 | -34.6679 | -21.1431
12.6236 | -34.6683 | -21.1515
12.6297 | -34.6688 | -21.1603
12.6358 | -34.6693 | -21.1688
12.6419 | -34.6698 | -21.1778
12.648 | -34.6703 | -21.1869
12.6541 | -34.6707 | -21.1956
12.6602 | -34.6712 | -21.204
12.6663 | -34.6716 | -21.2126
12.6724 | -34.6721 | -21.2214
12.6785 | -34.6726 | -21.2304
12.6847 | -34.673 | -21.2393
12.6908 | -34.6735 | -21.2479
12.6969 | -34.674 | -21.2569
12.703 | -34.6744 | -21.2657
12.7091 | -34.6749 | -21.2745
12.7152 | -34.6753 | -21.2832
12.7213 | -34.6758 | -21.2921
12.7274 | -34.6763 | -21.3009
12.7335 | -34.6767 | -21.3091
12.7396 | -34.6772 | -21.3174
12.7457 | -34.6776 | -21.3263
12.7518 | -34.6781 | -21.335
12.7579 | -34.6785 | -21.3441
12.764 | -34.679 | -21.3529
12.7701 | -34.6794 | -21.3614
12.7762 | -34.6798 | -21.3703
12.7823 | -34.6803 | -21.379
12.7884 | -34.6807 | -21.3876
12.7945 | -34.6812 | -21.396
12.8006 | -34.6817 | -21.4041
12.8067 | -34.6821 | -21.4132
12.8128 | -34.6825 | -21.4216
12.8189 | -34.6829 | -21.4299
12.8251 | -34.6833 | -21.4385
12.8312 | -34.6838 | -21.4472
12.8373 | -34.6842 | -21.4559
12.8434 | -34.6847 | -21.4645
12.8495 | -34.6851 | -21.4731
12.8556 | -34.6855 | -21.4819
12.8617 | -34.686 | -21.4903
12.8678 | -34.6864 | -21.499
12.8739 | -34.6869 | -21.5073
12.88 | -34.6873 | -21.5161
12.8861 | -34.6877 | -21.5248
12.8922 | -34.6882 | -21.5336
12.8983 | -34.6887 | -21.542
12.9044 | -34.6892 | -21.5506
12.9105 | -34.6897 | -21.5594
12.9166 | -34.6901 | -21.5684
12.9227 | -34.6906 | -21.5776
12.9288 | -34.6911 | -21.5864
12.9349 | -34.6916 | -21.5954
12.941 | -34.6921 | -21.6044
12.9471 | -34.6926 | -21.6132
12.9532 | -34.6931 | -21.6217
12.9593 | -34.6935 | -21.6307
12.9654 | -34.6939 | -21.6395
12.9716 | -34.6943 | -21.6483
12.9777 | -34.6948 | -21.6566
12.9838 | -34.6952 | -21.6651
12.9899 | -34.6956 | -21.674
12.996 | -34.6961 | -21.6828
13.0021 | -34.6965 | -21.6915
13.0082 | -34.6969 | -21.7001
13.0143 | -34.6974 | -21.7086
13.0204 | -34.6978 | -21.7171
13.0265 | -34.6983 | -21.7255
13.0326 | -34.6987 | -21.7344
13.0387 | -34.6992 | -21.7434
13.0448 | -34.6996 | -21.7519
13.0509 | -34.7001 | -21.7605
13.057 | -34.7006 | -21.7689
13.0631 | -34.7011 | -21.7774
13.0692 | -34.7016 | -21.7858
13.0753 | -34.7021 | -21.7945
13.0814 | -34.7025 | -21.8029
13.0875 | -34.703 | -21.8112
13.0936 | -34.7034 | -21.8198
13.0997 | -34.7039 | -21.828
13.1058 | -34.7043 | -21.836
13.112 | -34.7047 | -21.8448
13.1181 | -34.7052 | -21.8535
13.1242 | -34.7056 | -21.8622
13.1303 | -34.7061 | -21.8712
13.1364 | -34.7065 | -21.8798
13.1425 | -34.707 | -21.8882
13.1486 | -34.7074 | -21.8967
13.1547 | -34.7079 | -21.9051
13.1608 | -34.7084 | -21.9135
13.1669 | -34.7088 | -21.9221
13.173 | -34.7092 | -21.9307
13.1791 | -34.7096 | -21.9393
13.1852 | -34.7101 | -21.9483
13.1913 | -34.7106 | -21.9571
13.1974 | -34.711 | -21.9658
13.2035 | -34.7114 | -21.9744
13.2096 | -34.7119 | -21.983
13.2157 | -34.7124 | -21.9917
13.2218 | -34.7129 | -22.0004
13.2279 | -34.7133 | -22.0089
13.234 | -34.7138 | -22.0175
13.2401 | -34.7142 | -22.026
13.2462 | -34.7147 | -22.0346
13.2524 | -34.7151 | -22.0431
13.2585 | -34.7155 | -22.0521
13.2646 | -34.716 | -22.0608
13.2707 | -34.7164 | -22.0695
13.2768 | -34.7169 | -22.0784
13.2829 | -34.7173 | -22.087
13.289 | -34.7177 | -22.0956
13.2951 | -34.7182 | -22.104
13.3012 | -34.7186 | -22.1125
13.3073 | -34.719 | -22.1214
13.3134 | -34.7195 | -22.13
13.3195 | -34.7199 | -22.1385
13.3256 | -34.7204 | -22.1472
13.3317 | -34.7208 | -22.1555
13.3378 | -34.7212 | -22.1647
13.3439 | -34.7217 | -22.1733
13.35 | -34.7221 | -22.1818
13.3561 | -34.7226 | -22.1904
13.3622 | -34.723 | -22.1989
13.3683 | -34.7234 | -22.2073
13.3744 | -34.7238 | -22.2157
13.3805 | -34.7242 | -22.2245
13.3866 | -34.7246 | -22.2328
13.3927 | -34.7251 | -22.2413
13.3989 | -34.7255 | -22.2497
13.405 | -34.726 | -22.2581
13.4111 | -34.7265 | -22.2668
13.4172 | -34.7269 | -22.2753
13.4233 | -34.7273 | -22.2842
13.4294 | -34.7278 | -22.2925
13.4355 | -34.7282 | -22.3019
13.4416 | -34.7286 | -22.3104
13.4477 | -34.729 | -22.3192
13.4538 | -34.7295 | -22.3278
13.4599 | -34.73 | -22.3366
13.466 | -34.7304 | -22.3454
13.4721 | -34.7308 | -22.3542
13.4782 | -34.7313 | -22.3624
13.4843 | -34.7319 | -22.371
13.4904 | -34.7323 | -22.3789
13.4965 | -34.7327 | -22.3877
13.5026 | -34.7331 | -22.3962
13.5087 | -34.7336 | -22.4046
13.5148 | -34.734 | -22.4134
13.5209 | -34.7344 | -22.4217
13.527 | -34.7349 | -22.4305
13.5331 | -34.7353 | -22.4389
13.5393 | -34.7357 | -22.4474
13.5454 | -34.7361 | -22.4557
13.5515 | -34.7366 | -22.4643
13.5576 | -34.737 | -22.4731
13.5637 | -34.7375 | -22.4816
13.5698 | -34.738 | -22.4903
13.5759 | -34.7383 | -22.4984
13.582 | -34.7387 | -22.507
13.5881 | -34.7391 | -22.5157
13.5942 | -34.7396 | -22.5243
13.6003 | -34.74 | -22.533
13.6064 | -34.7404 | -22.5419
13.6125 | -34.7409 | -22.5507
13.6186 | -34.7414 | -22.5592
13.6247 | -34.7418 | -22.5676
13.6308 | -34.7422 | -22.5758
13.6369 | -34.7427 | -22.5843
13.643 | -34.7431 | -22.5929
13.6491 | -34.7435 | -22.6017
13.6552 | -34.744 | -22.6102
13.6613 | -34.7444 | -22.6189
13.6674 | -34.7448 | -22.6276
13.6735 | -34.7453 | -22.6365
13.6796 | -34.7457 | -22.6447
13.6858 | -34.7461 | -22.6531
13.6919 | -34.7466 | -22.6616
13.698 | -34.747 | -22.6701
13.7041 | -34.7474 | -22.6789
13.7102 | -34.7478 | -22.6876
13.7163 | -34.7484 | -22.6963
13.7224 | -34.7488 | -22.7049
13.7285 | -34.7493 | -22.7134
13.7346 | -34.7497 | -22.722
13.7407 | -34.7501 | -22.7307
13.7468 | -34.7505 | -22.7393
13.7529 | -34.7509 | -22.7477
13.759 | -34.7514 | -22.7563
13.7651 | -34.7518 | -22.765
13.7712 | -34.7522 | -22.7738
13.7773 | -34.7526 | -22.7824
13.7834 | -34.7531 | -22.791
13.7895 | -34.7536 | -22.7995
13.7956 | -34.754 | -22.8083
13.8017 | -34.7544 | -22.8167
13.8078 | -34.7549 | -22.8255
13.8139 | -34.7554 | -22.834
13.82 | -34.7558 | -22.8422
13.8262 | -34.7562 | -22.8512
13.8323 | -34.7567 | -22.86
13.8384 | -34.7571 | -22.8686
13.8445 | -34.7575 | -22.8772
13.8506 | -34.758 | -22.8857
13.8567 | -34.7584 | -22.8947
13.8628 | -34.7588 | -22.9033
13.8689 | -34.7592 | -22.9118
13.875 | -34.7597 | -22.9202
13.8811 | -34.7602 | -22.9286
13.8872 | -34.7606 | -22.9371
13.8933 | -34.761 | -22.9457
13.8994 | -34.7615 | -22.954
13.9055 | -34.7619 | -22.9626
13.9116 | -34.7624 | -22.9709
13.9177 | -34.7628 | -22.979
13.9238 | -34.7633 | -22.9877
13.9299 | -34.7638 | -22.9962
13.936 | -34.7642 | -23.0047
13.9421 | -34.7646 | -23.0133
13.9482 | -34.7651 | -23.0218
13.9543 | -34.7655 | -23.03
13.9604 | -34.7659 | -23.0386
13.9665 | -34.7664 | -23.0471
13.9727 | -34.7668 | -23.0555
13.9788 | -34.7672 | -23.0642
13.9849 | -34.7676 | -23.0728
13.991 | -34.768 | -23.0814
13.9971 | -34.7685 | -23.09
14.0032 | -34.7689 | -23.0987
14.0093 | -34.7693 | -23.107
14.0154 | -34.7697 | -23.1159
14.0215 | -34.7701 | -23.1241
14.0276 | -34.7705 | -23.1326
14.0337 | -34.771 | -23.1413
14.0398 | -34.7715 | -23.1497
14.0459 | -34.7719 | -23.1583
14.052 | -34.7723 | -23.1668
14.0581 | -34.7727 | -23.1748
14.0642 | -34.7732 | -23.1835
14.0703 | -34.7736 | -23.1917
14.0764 | -34.7741 | -23.2006
14.0825 | -34.7745 | -23.2091
14.0886 | -34.7749 | -23.2174
14.0947 | -34.7753 | -23.2257
14.1008 | -34.7758 | -23.2344
14.1069 | -34.7762 | -23.2428
14.1131 | -34.7766 | -23.2516
14.1192 | -34.777 | -23.2595
14.1253 | -34.7775 | -23.2678
14.1314 | -34.7779 | -23.2764
14.1375 | -34.7784 | -23.285
14.1436 | -34.7789 | -23.2935
14.1497 | -34.7793 | -23.3018
14.1558 | -34.7797 | -23.3105
14.1619 | -34.7802 | -23.319
14.168 | -34.7806 | -23.3273
14.1741 | -34.7811 | -23.3355
14.1802 | -34.7815 | -23.3438
14.1863 | -34.782 | -23.3522
14.1924 | -34.7824 | -23.3606
14.1985 | -34.7828 | -23.3689
14.2046 | -34.7833 | -23.3774
14.2107 | -34.7837 | -23.3853
14.2168 | -34.7842 | -23.394
14.2229 | -34.7846 | -23.4026
14.229 | -34.785 | -23.4112
14.2351 | -34.7855 | -23.4194
14.2412 | -34.7859 | -23.4275
14.2473 | -34.7864 | -23.436
14.2534 | -34.7868 | -23.4445
14.2596 | -34.7872 | -23.4528
14.2657 | -34.7877 | -23.4614
14.2718 | -34.7881 | -23.47
14.2779 | -34.7886 | -23.4781
14.284 | -34.789 | -23.4865
14.2901 | -34.7895 | -23.4946
14.2962 | -34.7899 | -23.5028
14.3023 | -34.7904 | -23.5112
14.3084 | -34.7908 | -23.5199
14.3145 | -34.7913 | -23.5281
14.3206 | -34.7916 | -23.5369
14.3267 | -34.7921 | -23.5453
14.3328 | -34.7925 | -23.5536
14.3389 | -34.793 | -23.5617
14.345 | -34.7934 | -23.5702
14.3511 | -34.7938 | -23.5782
14.3572 | -34.7942 | -23.5867
14.3633 | -34.7947 | -23.5951
14.3694 | -34.7951 | -23.6033
14.3755 | -34.7955 | -23.6113
14.3816 | -34.796 | -23.6196
14.3877 | -34.7964 | -23.6277
14.3938 | -34.7969 | -23.6363
14.4 | -34.7973 | -23.6444
14.4061 | -34.7977 | -23.6528
14.4122 | -34.7981 | -23.6613
14.4183 | -34.7985 | -23.6696
14.4244 | -34.7989 | -23.6777
14.4305 | -34.7994 | -23.6864
14.4366 | -34.7998 | -23.6948
14.4427 | -34.8003 | -23.7033
14.4488 | -34.8007 | -23.7113
14.4549 | -34.8011 | -23.7193
14.461 | -34.8016 | -23.7276
14.4671 | -34.802 | -23.736
14.4732 | -34.8025 | -23.7439
14.4793 | -34.8029 | -23.753
14.4854 | -34.8035 | -23.7609
14.4915 | -34.8039 | -23.7689
14.4976 | -34.8044 | -23.7775
14.5037 | -34.8048 | -23.7859
14.5098 | -34.8053 | -23.7941
14.5159 | -34.8057 | -23.8024
14.522 | -34.8062 | -23.8109
14.5281 | -34.8066 | -23.8188
14.5342 | -34.807 | -23.8273
14.5403 | -34.8075 | -23.8357
14.5465 | -34.808 | -23.8442
14.5526 | -34.8084 | -23.8528
14.5587 | -34.8088 | -23.861
14.5648 | -34.8093 | -23.869
14.5709 | -34.8097 | -23.8775
14.577 | -34.8102 | -23.8859
14.5831 | -34.8106 | -23.8943
14.5892 | -34.8111 | -23.9022
14.5953 | -34.8116 | -23.9103
14.6014 | -34.812 | -23.9185
14.6075 | -34.8125 | -23.9265
14.6136 | -34.8129 | -23.9346
14.6197 | -34.8133 | -23.943
14.6258 | -34.8138 | -23.9511
14.6319 | -34.8143 | -23.9591
14.638 | -34.8147 | -23.9674
14.6441 | -34.8152 | -23.9752
14.6502 | -34.8156 | -23.9834
14.6563 | -34.816 | -23.9917
14.6624 | -34.8165 | -23.9998
14.6685 | -34.817 | -24.0078
14.6746 | -34.8174 | -24.0161
14.6807 | -34.8179 | -24.0243
14.6869 | -34.8183 | -24.0324
14.693 | -34.8188 | -24.0404
14.6991 | -34.8192 | -24.0486
14.7052 | -34.8197 | -24.0566
14.7113 | -34.8201 | -24.0651
14.7174 | -34.8206 | -24.0734
14.7235 | -34.821 | -24.0814
14.7296 | -34.8215 | -24.0896
14.7357 | -34.822 | -24.0976
14.7418 | -34.8224 | -24.1057
14.7479 | -34.8228 | -24.1139
14.754 | -34.8233 | -24.1221
14.7601 | -34.8237 | -24.1304
14.7662 | -34.8242 | -24.1385
14.7723 | -34.8246 | -24.1465
14.7784 | -34.8251 | -24.1548
14.7845 | -34.8255 | -24.1631
14.7906 | -34.826 | -24.1715
14.7967 | -34.8264 | -24.1796
14.8028 | -34.8268 | -24.1875
14.8089 | -34.8273 | -24.1956
14.815 | -34.8277 | -24.2035
14.8211 | -34.8281 | -24.2116
14.8272 | -34.8286 | -24.2198
14.8334 | -34.829 | -24.2278
14.8395 | -34.8295 | -24.2358
14.8456 | -34.83 | -24.2437
14.8517 | -34.8304 | -24.2519
14.8578 | -34.8309 | -24.2601
14.8639 | -34.8313 | -24.2684
14.87 | -34.8318 | -24.2764
14.8761 | -34.8323 | -24.284
14.8822 | -34.8327 | -24.2921
14.8883 | -34.8331 | -24.3001
14.8944 | -34.8336 | -24.3081
14.9005 | -34.8341 | -24.316
14.9066 | -34.8345 | -24.3241
14.9127 | -34.8349 | -24.332
14.9188 | -34.8354 | -24.3401
14.9249 | -34.8358 | -24.348
14.931 | -34.8363 | -24.3559
14.9371 | -34.8367 | -24.3637
14.9432 | -34.8372 | -24.3719
14.9493 | -34.8377 | -24.3799
14.9554 | -34.8381 | -24.3881
14.9615 | -34.8386 | -24.3963
14.9676 | -34.839 | -24.4047
14.9738 | -34.8395 | -24.4127
14.9799 | -34.84 | -24.4211
14.986 | -34.8404 | -24.4286
14.9921 | -34.841 | -24.4366
14.9982 | -34.8414 | -24.4449
15.0043 | -34.8418 | -24.4529
15.0104 | -34.8423 | -24.4609
15.0165 | -34.8428 | -24.4691
15.0226 | -34.8433 | -24.4769
15.0287 | -34.8437 | -24.4853
15.0348 | -34.8442 | -24.4934
15.0409 | -34.8446 | -24.502
15.047 | -34.8451 | -24.5104
15.0531 | -34.8457 | -24.5184
15.0592 | -34.8461 | -24.5263
15.0653 | -34.8466 | -24.5343
15.0714 | -34.847 | -24.5421
15.0775 | -34.8475 | -24.5496
15.0836 | -34.8479 | -24.5577
15.0897 | -34.8485 | -24.5654
15.0958 | -34.8489 | -24.5733
15.1019 | -34.8494 | -24.5815
15.108 | -34.8498 | -24.5897
15.1141 | -34.8502 | -24.5977
15.1203 | -34.8508 | -24.6057
15.1264 | -34.8513 | -24.6137
15.1325 | -34.8518 | -24.6217
15.1386 | -34.8523 | -24.6299
15.1447 | -34.8529 | -24.6376
15.1508 | -34.8534 | -24.6459
15.1569 | -34.8538 | -24.6542
15.163 | -34.8544 | -24.6616
15.1691 | -34.8549 | -24.6694
15.1752 | -34.8554 | -24.6773
15.1813 | -34.8559 | -24.6853
15.1874 | -34.8563 | -24.6934
15.1935 | -34.8569 | -24.7011
15.1996 | -34.8574 | -24.7091
15.2057 | -34.8578 | -24.7169
15.2118 | -34.8583 | -24.7249
15.2179 | -34.8588 | -24.7327
15.224 | -34.8593 | -24.7407
15.2301 | -34.8598 | -24.7481
15.2362 | -34.8603 | -24.7557
15.2423 | -34.8608 | -24.7641
15.2484 | -34.8612 | -24.772
15.2545 | -34.8617 | -24.7799
15.2607 | -34.8622 | -24.7879
15.2668 | -34.8628 | -24.7957
15.2729 | -34.8633 | -24.8034
15.279 | -34.8638 | -24.8114
15.2851 | -34.8643 | -24.8195
15.2912 | -34.8647 | -24.8274
15.2973 | -34.8652 | -24.8353
15.3034 | -34.8658 | -24.8429
15.3095 | -34.8662 | -24.8506
15.3156 | -34.8667 | -24.8587
15.3217 | -34.8672 | -24.8661
15.3278 | -34.8678 | -24.874
15.3339 | -34.8683 | -24.8818
15.34 | -34.8688 | -24.8895
15.3461 | -34.8693 | -24.8974
15.3522 | -34.8698 | -24.9051
15.3583 | -34.8703 | -24.913
15.3644 | -34.8708 | -24.9209
15.3705 | -34.8713 | -24.9285
15.3766 | -34.8718 | -24.9364
15.3827 | -34.8723 | -24.9441
15.3888 | -34.8728 | -24.9518
15.3949 | -34.8733 | -24.9595
15.401 | -34.8738 | -24.9677
15.4072 | -34.8743 | -24.9751
15.4133 | -34.8748 | -24.9828
15.4194 | -34.8753 | -24.9905
15.4255 | -34.8758 | -24.9984
15.4316 | -34.8763 | -25.0062
15.4377 | -34.8768 | -25.0141
15.4438 | -34.8774 | -25.0216
15.4499 | -34.8778 | -25.0293
15.456 | -34.8783 | -25.0371
15.4621 | -34.8789 | -25.045
15.4682 | -34.8794 | -25.0531
15.4743 | -34.88 | -25.0609
15.4804 | -34.8804 | -25.0688
15.4865 | -34.8809 | -25.0765
15.4926 | -34.8814 | -25.0847
15.4987 | -34.8819 | -25.0927
15.5048 | -34.8824 | -25.1003
15.5109 | -34.883 | -25.1079
15.517 | -34.8835 | -25.1157
15.5231 | -34.884 | -25.1236
15.5292 | -34.8845 | -25.1311
15.5353 | -34.885 | -25.1388
15.5414 | -34.8855 | -25.1467
15.5476 | -34.886 | -25.1547
15.5537 | -34.8865 | -25.1624
15.5598 | -34.8871 | -25.1701
15.5659 | -34.8876 | -25.1781
15.572 | -34.8881 | -25.1856
15.5781 | -34.8886 | -25.1933
15.5842 | -34.8891 | -25.2009
15.5903 | -34.8896 | -25.2084
15.5964 | -34.8902 | -25.2159
15.6025 | -34.8907 | -25.2235
15.6086 | -34.8912 | -25.2315
15.6147 | -34.8917 | -25.2394
15.6208 | -34.8923 | -25.2467
15.6269 | -34.8928 | -25.2543
15.633 | -34.8933 | -25.2618
15.6391 | -34.8938 | -25.2696
15.6452 | -34.8944 | -25.2771
15.6513 | -34.8949 | -25.285
15.6574 | -34.8954 | -25.2922
15.6635 | -34.8959 | -25.3001
15.6696 | -34.8964 | -25.308
15.6757 | -34.897 | -25.3159
15.6818 | -34.8974 | -25.3239
15.688 | -34.898 | -25.3313
15.6941 | -34.8985 | -25.3389
15.7002 | -34.899 | -25.3465
15.7063 | -34.8995 | -25.3545
15.7124 | -34.9001 | -25.3624
15.7185 | -34.9006 | -25.3701
15.7246 | -34.9012 | -25.3776
15.7307 | -34.9017 | -25.3856
15.7368 | -34.9022 | -25.3936
15.7429 | -34.9027 | -25.4013
15.749 | -34.9032 | -25.4089
15.7551 | -34.9038 | -25.4169
15.7612 | -34.9043 | -25.424
15.7673 | -34.9048 | -25.4315
15.7734 | -34.9053 | -25.439
15.7795 | -34.9059 | -25.4471
15.7856 | -34.9064 | -25.4544
15.7917 | -34.9069 | -25.4619
15.7978 | -34.9075 | -25.4696
15.8039 | -34.908 | -25.4771
15.81 | -34.9084 | -25.4847
15.8161 | -34.909 | -25.4924
15.8222 | -34.9095 | -25.4999
15.8283 | -34.9101 | -25.5074
15.8345 | -34.9106 | -25.5149
15.8406 | -34.9111 | -25.5226
15.8467 | -34.9116 | -25.5304
15.8528 | -34.9122 | -25.5379
15.8589 | -34.9127 | -25.5454
15.865 | -34.9132 | -25.553
15.8711 | -34.9138 | -25.5612
15.8772 | -34.9144 | -25.569
15.8833 | -34.9149 | -25.577
15.8894 | -34.9154 | -25.5846
15.8955 | -34.916 | -25.5923
15.9016 | -34.9165 | -25.6
15.9077 | -34.9171 | -25.6079
15.9138 | -34.9176 | -25.6153
15.9199 | -34.9181 | -25.6232
15.926 | -34.9186 | -25.6311
15.9321 | -34.9193 | -25.6385
15.9382 | -34.9198 | -25.6457
15.9443 | -34.9204 | -25.6538
15.9504 | -34.921 | -25.6615
15.9565 | -34.9215 | -25.6692
15.9626 | -34.9221 | -25.677
15.9687 | -34.9226 | -25.6845
15.9749 | -34.9232 | -25.6923
15.981 | -34.9237 | -25.7
15.9871 | -34.9242 | -25.7078
15.9932 | -34.9248 | -25.7157
15.9993 | -34.9253 | -25.7232
16.0054 | -34.9259 | -25.7306
16.0115 | -34.9265 | -25.738
16.0176 | -34.927 | -25.746
16.0237 | -34.9276 | -25.7539
16.0298 | -34.9281 | -25.7609
16.0359 | -34.9287 | -25.7686
16.042 | -34.9292 | -25.7763
16.0481 | -34.9298 | -25.7839
16.0542 | -34.9303 | -25.7919
16.0603 | -34.9309 | -25.7995
16.0664 | -34.9314 | -25.807
16.0725 | -34.932 | -25.8145
16.0786 | -34.9325 | -25.8223
16.0847 | -34.933 | -25.8298
16.0908 | -34.9336 | -25.8374
16.0969 | -34.9341 | -25.8449
16.103 | -34.9346 | -25.8527
16.1091 | -34.9352 | -25.8602
16.1152 | -34.9357 | -25.8679
16.1214 | -34.9362 | -25.8755
16.1275 | -34.9368 | -25.8831
16.1336 | -34.9374 | -25.8908
16.1397 | -34.9379 | -25.8984
16.1458 | -34.9385 | -25.9059
16.1519 | -34.9391 | -25.9137
16.158 | -34.9396 | -25.9216
16.1641 | -34.9401 | -25.9294
16.1702 | -34.9406 | -25.9373
16.1763 | -34.9412 | -25.945
16.1824 | -34.9418 | -25.9524
16.1885 | -34.9424 | -25.96
16.1946 | -34.9429 | -25.9678
16.2007 | -34.9435 | -25.9757
16.2068 | -34.944 | -25.9831
16.2129 | -34.9445 | -25.9908
16.219 | -34.945 | -25.9984
16.2251 | -34.9456 | -26.0062
16.2312 | -34.9461 | -26.0139
16.2373 | -34.9467 | -26.0214
16.2434 | -34.9472 | -26.0291
16.2495 | -34.9478 | -26.0366
16.2556 | -34.9484 | -26.0443
16.2618 | -34.949 | -26.0518
16.2679 | -34.9495 | -26.0594
16.274 | -34.9501 | -26.0671
16.2801 | -34.9506 | -26.0746
16.2862 | -34.9512 | -26.0822
16.2923 | -34.9518 | -26.0899
16.2984 | -34.9523 | -26.0977
16.3045 | -34.9529 | -26.1051
16.3106 | -34.9535 | -26.1123
16.3167 | -34.954 | -26.1202
16.3228 | -34.9546 | -26.1281
16.3289 | -34.9551 | -26.1357
16.335 | -34.9556 | -26.1435
16.3411 | -34.9562 | -26.1511
16.3472 | -34.9567 | -26.1587
16.3533 | -34.9573 | -26.1664
16.3594 | -34.9578 | -26.1735
16.3655 | -34.9583 | -26.1812
16.3716 | -34.9589 | -26.1888
16.3777 | -34.9595 | -26.1963
16.3838 | -34.96 | -26.2038
16.3899 | -34.9606 | -26.2113
16.396 | -34.9611 | -26.2189
16.4021 | -34.9616 | -26.2264
16.4083 | -34.9622 | -26.2338
16.4144 | -34.9628 | -26.2411
16.4205 | -34.9633 | -26.2487
16.4266 | -34.9639 | -26.2564
16.4327 | -34.9644 | -26.264
16.4388 | -34.965 | -26.2715
16.4449 | -34.9655 | -26.279
16.451 | -34.9661 | -26.287
16.4571 | -34.9666 | -26.295
16.4632 | -34.9671 | -26.3025
16.4693 | -34.9676 | -26.3101
16.4754 | -34.9682 | -26.3177
16.4815 | -34.9688 | -26.3256
16.4876 | -34.9693 | -26.3333
16.4937 | -34.9699 | -26.3407
16.4998 | -34.9705 | -26.3484
16.5059 | -34.9711 | -26.3559
16.512 | -34.9716 | -26.3637
16.5181 | -34.9722 | -26.3716
16.5242 | -34.9727 | -26.3793
16.5303 | -34.9732 | -26.387
16.5364 | -34.9738 | -26.3942
16.5425 | -34.9743 | -26.4019
16.5487 | -34.9748 | -26.4096
16.5548 | -34.9753 | -26.4172
16.5609 | -34.9759 | -26.4249
16.567 | -34.9764 | -26.4325
16.5731 | -34.9769 | -26.4404
16.5792 | -34.9775 | -26.4479
16.5853 | -34.9781 | -26.4554
16.5914 | -34.9787 | -26.4633
16.5975 | -34.9792 | -26.4712
16.6036 | -34.9797 | -26.4788
16.6097 | -34.9803 | -26.4865
16.6158 | -34.9808 | -26.4945
16.6219 | -34.9813 | -26.5022
16.628 | -34.9818 | -26.5099
16.6341 | -34.9824 | -26.5175
16.6402 | -34.9829 | -26.5248
16.6463 | -34.9834 | -26.5324
16.6524 | -34.9839 | -26.5401
16.6585 | -34.9845 | -26.5478
16.6646 | -34.9851 | -26.5554
16.6707 | -34.9856 | -26.5629
16.6768 | -34.9862 | -26.5704
16.6829 | -34.9867 | -26.5781
16.689 | -34.9872 | -26.586
16.6952 | -34.9878 | -26.5936
16.7013 | -34.9883 | -26.6014
16.7074 | -34.9889 | -26.6089
16.7135 | -34.9894 | -26.6164
16.7196 | -34.99 | -26.6238
16.7257 | -34.9905 | -26.6311
16.7318 | -34.9911 | -26.6387
16.7379 | -34.9916 | -26.6465
16.744 | -34.9922 | -26.6535
16.7501 | -34.9927 | -26.661
16.7562 | -34.9933 | -26.6689
16.7623 | -34.9938 | -26.6764
16.7684 | -34.9943 | -26.6842
16.7745 | -34.9949 | -26.6913
16.7806 | -34.9954 | -26.6991
16.7867 | -34.996 | -26.7068
16.7928 | -34.9965 | -26.7145
16.7989 | -34.997 | -26.722
16.805 | -34.9975 | -26.7296
16.8111 | -34.998 | -26.7375
16.8172 | -34.9985 | -26.7452
16.8233 | -34.9991 | -26.753
16.8294 | -34.9997 | -26.7607
16.8356 | -35.0002 | -26.7687
16.8417 | -35.0008 | -26.7764
16.8478 | -35.0013 | -26.7836
16.8539 | -35.0018 | -26.7909
16.86 | -35.0024 | -26.7985
16.8661 | -35.0029 | -26.806
16.8722 | -35.0035 | -26.8135
16.8783 | -35.004 | -26.8208
16.8844 | -35.0046 | -26.8284
16.8905 | -35.0051 | -26.8362
16.8966 | -35.0056 | -26.844
16.9027 | -35.0062 | -26.8516
16.9088 | -35.0067 | -26.8597
16.9149 | -35.0073 | -26.8673
16.921 | -35.0077 | -26.8751
16.9271 | -35.0083 | -26.8832
16.9332 | -35.0088 | -26.8909
16.9393 | -35.0093 | -26.8987
16.9454 | -35.0099 | -26.9061
16.9515 | -35.0104 | -26.9142
16.9576 | -35.0109 | -26.9221
16.9637 | -35.0115 | -26.9295
16.9698 | -35.0121 | -26.9367
16.9759 | -35.0126 | -26.9443
16.9821 | -35.0132 | -26.9523
16.9882 | -35.0137 | -26.9602
16.9943 | -35.0143 | -26.9678
17.0004 | -35.0148 | -26.9755
17.0065 | -35.0153 | -26.9833
17.0126 | -35.0158 | -26.991
17.0187 | -35.0163 | -26.9986
17.0248 | -35.0169 | -27.0061
17.0309 | -35.0174 | -27.0137
17.037 | -35.0179 | -27.0213
17.0431 | -35.0185 | -27.0286
17.0492 | -35.019 | -27.0361
17.0553 | -35.0194 | -27.0438
17.0614 | -35.0199 | -27.0511
17.0675 | -35.0204 | -27.0587
17.0736 | -35.021 | -27.0664
17.0797 | -35.0215 | -27.0734
17.0858 | -35.022 | -27.0812
17.0919 | -35.0225 | -27.0888
17.098 | -35.023 | -27.0962
17.1041 | -35.0235 | -27.1034
17.1102 | -35.0241 | -27.1111
17.1163 | -35.0247 | -27.1188
17.1225 | -35.0252 | -27.1262
17.1286 | -35.0257 | -27.1334
17.1347 | -35.0262 | -27.1408
17.1408 | -35.0267 | -27.1486
17.1469 | -35.0273 | -27.1561
17.153 | -35.0278 | -27.1637
17.1591 | -35.0283 | -27.1709
17.1652 | -35.0288 | -27.1783
17.1713 | -35.0293 | -27.1858
17.1774 | -35.0299 | -27.1935
17.1835 | -35.0304 | -27.2014
17.1896 | -35.0309 | -27.2089
17.1957 | -35.0314 | -27.2165
17.2018 | -35.0319 | -27.2245
17.2079 | -35.0324 | -27.232
17.214 | -35.0329 | -27.2395
17.2201 | -35.0334 | -27.2472
17.2262 | -35.0339 | -27.255
17.2323 | -35.0345 | -27.2626
17.2384 | -35.035 | -27.2703
17.2445 | -35.0355 | -27.2777
17.2506 | -35.036 | -27.2856
17.2567 | -35.0365 | -27.2933
17.2628 | -35.037 | -27.3012
17.269 | -35.0375 | -27.309
17.2751 | -35.0381 | -27.3165
17.2812 | -35.0386 | -27.324
17.2873 | -35.0391 | -27.3318
17.2934 | -35.0396 | -27.3395
17.2995 | -35.0401 | -27.3468
17.3056 | -35.0406 | -27.3544
17.3117 | -35.0411 | -27.362
17.3178 | -35.0416 | -27.3696
17.3239 | -35.0421 | -27.3771
17.33 | -35.0426 | -27.3844
17.3361 | -35.0431 | -27.392
17.3422 | -35.0436 | -27.3995
17.3483 | -35.0441 | -27.4071
17.3544 | -35.0446 | -27.4148
17.3605 | -35.0452 | -27.4225
17.3666 | -35.0457 | -27.4302
17.3727 | -35.0462 | -27.4375
17.3788 | -35.0467 | -27.4451
17.3849 | -35.0472 | -27.4527
17.391 | -35.0477 | -27.4601
17.3971 | -35.0483 | -27.4677
17.4032 | -35.0489 | -27.4756
17.4094 | -35.0493 | -27.4831
17.4155 | -35.0498 | -27.4907
17.4216 | -35.0503 | -27.4983
17.4277 | -35.0507 | -27.5061
17.4338 | -35.0513 | -27.5136
17.4399 | -35.0518 | -27.5213
17.446 | -35.0523 | -27.5289
17.4521 | -35.0528 | -27.5361
17.4582 | -35.0533 | -27.5436
17.4643 | -35.0538 | -27.5512
17.4704 | -35.0543 | -27.5586
17.4765 | -35.0548 | -27.5661
17.4826 | -35.0553 | -27.5736
17.4887 | -35.0558 | -27.5814
17.4948 | -35.0563 | -27.589
17.5009 | -35.0568 | -27.5964
17.507 | -35.0573 | -27.6041
17.5131 | -35.0578 | -27.6115
17.5192 | -35.0583 | -27.6189
17.5253 | -35.0587 | -27.6266
17.5314 | -35.0592 | -27.6342
17.5375 | -35.0598 | -27.6417
17.5436 | -35.0602 | -27.6496
17.5497 | -35.0607 | -27.6576
17.5559 | -35.0612 | -27.665
17.562 | -35.0617 | -27.6724
17.5681 | -35.0622 | -27.6799
17.5742 | -35.0627 | -27.6877
17.5803 | -35.0632 | -27.6951
17.5864 | -35.0638 | -27.7025
17.5925 | -35.0643 | -27.7103
17.5986 | -35.0648 | -27.7178
17.6047 | -35.0653 | -27.7257
17.6108 | -35.0658 | -27.733
17.6169 | -35.0663 | -27.7405
17.623 | -35.0668 | -27.7479
17.6291 | -35.0673 | -27.7554
17.6352 | -35.0677 | -27.7628
17.6413 | -35.0682 | -27.7706
17.6474 | -35.0687 | -27.7782
17.6535 | -35.0692 | -27.786
17.6596 | -35.0697 | -27.7935
17.6657 | -35.0702 | -27.8008
17.6718 | -35.0707 | -27.8083
17.6779 | -35.0711 | -27.8157
17.684 | -35.0717 | -27.8232
17.6901 | -35.0722 | -27.8307
17.6963 | -35.0726 | -27.8378
17.7024 | -35.0731 | -27.8452
17.7085 | -35.0736 | -27.8525
17.7146 | -35.0741 | -27.8601
17.7207 | -35.0746 | -27.8672
17.7268 | -35.0751 | -27.8748
17.7329 | -35.0756 | -27.8826
17.739 | -35.0761 | -27.8902
17.7451 | -35.0767 | -27.8978
17.7512 | -35.0771 | -27.9054
17.7573 | -35.0776 | -27.9129
17.7634 | -35.0781 | -27.9203
17.7695 | -35.0786 | -27.9276
17.7756 | -35.0791 | -27.935
17.7817 | -35.0796 | -27.9422
17.7878 | -35.0802 | -27.9497
17.7939 | -35.0806 | -27.9573
17.8 | -35.0811 | -27.965
17.8061 | -35.0816 | -27.9727
17.8122 | -35.0821 | -27.9801
17.8183 | -35.0826 | -27.9875
17.8244 | -35.0831 | -27.9948
17.8305 | -35.0836 | -28.0022
17.8366 | -35.0841 | -28.0093
17.8428 | -35.0845 | -28.0171
17.8489 | -35.085 | -28.0244
17.855 | -35.0855 | -28.0315
17.8611 | -35.086 | -28.0388
17.8672 | -35.0865 | -28.0461
17.8733 | -35.087 | -28.0539
17.8794 | -35.0875 | -28.0614
17.8855 | -35.088 | -28.0686
17.8916 | -35.0884 | -28.0759
17.8977 | -35.0889 | -28.0832
17.9038 | -35.0893 | -28.0905
17.9099 | -35.0898 | -28.0981
17.916 | -35.0903 | -28.1057
17.9221 | -35.0908 | -28.1133
17.9282 | -35.0912 | -28.1208
17.9343 | -35.0917 | -28.1284
17.9404 | -35.0922 | -28.1358
17.9465 | -35.0927 | -28.1435
17.9526 | -35.0931 | -28.1511
17.9587 | -35.0936 | -28.1585
17.9648 | -35.0941 | -28.1659
17.9709 | -35.0946 | -28.1733
17.977 | -35.095 | -28.1807
17.9832 | -35.0955 | -28.1881
17.9893 | -35.096 | -28.1957
17.9954 | -35.0965 | -28.2031
18.0015 | -35.0971 | -28.2106
18.0076 | -35.0974 | -28.2178
18.0137 | -35.0979 | -28.2256
18.0198 | -35.0984 | -28.233
18.0259 | -35.0988 | -28.2406
18.032 | -35.0993 | -28.2479
18.0381 | -35.0998 | -28.2555
18.0442 | -35.1003 | -28.2629
18.0503 | -35.1008 | -28.2703
18.0564 | -35.1012 | -28.2776
18.0625 | -35.1017 | -28.2852
18.0686 | -35.1022 | -28.2925
18.0747 | -35.1027 | -28.2999
18.0808 | -35.1032 | -28.3072
18.0869 | -35.1037 | -28.3147
18.093 | -35.1042 | -28.3222
18.0991 | -35.1046 | -28.3299
18.1052 | -35.1051 | -28.3375
18.1113 | -35.1056 | -28.3447
18.1174 | -35.106 | -28.3522
18.1236 | -35.1066 | -28.3596
18.1297 | -35.107 | -28.367
18.1358 | -35.1076 | -28.3746
18.1419 | -35.108 | -28.3822
18.148 | -35.1085 | -28.3899
18.1541 | -35.109 | -28.3972
18.1602 | -35.1095 | -28.4048
18.1663 | -35.1101 | -28.4123
18.1724 | -35.1105 | -28.4202
18.1785 | -35.111 | -28.4279
18.1846 | -35.1114 | -28.435
18.1907 | -35.1119 | -28.443
18.1968 | -35.1123 | -28.4506
18.2029 | -35.1128 | -28.4577
18.209 | -35.1133 | -28.4654
18.2151 | -35.1138 | -28.473
18.2212 | -35.1144 | -28.4799
18.2273 | -35.1149 | -28.4874
18.2334 | -35.1153 | -28.4948
18.2395 | -35.1157 | -28.5019
18.2456 | -35.1163 | -28.5096
18.2517 | -35.1167 | -28.517
18.2578 | -35.1172 | -28.5242
18.2639 | -35.1177 | -28.5317
18.2701 | -35.1182 | -28.5389
18.2762 | -35.1187 | -28.5466
18.2823 | -35.1192 | -28.5541
18.2884 | -35.1197 | -28.5618
18.2945 | -35.1202 | -28.5689
18.3006 | -35.1206 | -28.5765
18.3067 | -35.1212 | -28.5839
18.3128 | -35.1217 | -28.5915
18.3189 | -35.1222 | -28.5993
18.325 | -35.1227 | -28.6063
18.3311 | -35.1232 | -28.6141
18.3372 | -35.1236 | -28.6212
18.3433 | -35.1241 | -28.6283
18.3494 | -35.1245 | -28.6356
18.3555 | -35.125 | -28.6429
18.3616 | -35.1255 | -28.6499
18.3677 | -35.126 | -28.6572
18.3738 | -35.1265 | -28.6645
18.3799 | -35.127 | -28.6716
18.386 | -35.1274 | -28.6788
18.3921 | -35.1279 | -28.6863
18.3982 | -35.1284 | -28.6936
18.4043 | -35.1289 | -28.7006
18.4105 | -35.1293 | -28.708
18.4166 | -35.1298 | -28.7153
18.4227 | -35.1303 | -28.7228
18.4288 | -35.1308 | -28.7303
18.4349 | -35.1312 | -28.7377
18.441 | -35.1317 | -28.745
18.4471 | -35.1321 | -28.7526
18.4532 | -35.1326 | -28.7602
18.4593 | -35.1331 | -28.7675
18.4654 | -35.1336 | -28.7744
18.4715 | -35.1341 | -28.7819
18.4776 | -35.1346 | -28.7896
18.4837 | -35.135 | -28.7971
18.4898 | -35.1355 | -28.8044
18.4959 | -35.136 | -28.812
18.502 | -35.1366 | -28.8195
18.5081 | -35.1371 | -28.8269
18.5142 | -35.1374 | -28.8343
18.5203 | -35.1379 | -28.8413
18.5264 | -35.1383 | -28.8485
18.5325 | -35.1388 | -28.8561
18.5386 | -35.1393 | -28.8638
18.5447 | -35.1398 | -28.8711
18.5508 | -35.1403 | -28.8786
18.557 | -35.1407 | -28.886
18.5631 | -35.1412 | -28.8933
18.5692 | -35.1416 | -28.9006
18.5753 | -35.1421 | -28.9078
18.5814 | -35.1426 | -28.9152
18.5875 | -35.143 | -28.9229
18.5936 | -35.1435 | -28.9309
18.5997 | -35.144 | -28.9384
18.6058 | -35.1444 | -28.9458
18.6119 | -35.1449 | -28.9534
18.618 | -35.1454 | -28.9608
18.6241 | -35.1459 | -28.9684
18.6302 | -35.1464 | -28.976
18.6363 | -35.1468 | -28.983
18.6424 | -35.1473 | -28.9899
18.6485 | -35.1477 | -28.9973
18.6546 | -35.1482 | -29.0044
18.6607 | -35.1487 | -29.0117
18.6668 | -35.1492 | -29.0189
18.6729 | -35.1497 | -29.0264
18.679 | -35.1501 | -29.0338
18.6851 | -35.1506 | -29.0409
18.6912 | -35.1511 | -29.0483
18.6974 | -35.1515 | -29.0558
18.7035 | -35.152 | -29.0631
18.7096 | -35.1525 | -29.0705
18.7157 | -35.153 | -29.0779
18.7218 | -35.1535 | -29.085
18.7279 | -35.154 | -29.0921
18.734 | -35.1545 | -29.0993
18.7401 | -35.155 | -29.1065
18.7462 | -35.1554 | -29.1131
18.7523 | -35.1559 | -29.1203
18.7584 | -35.1564 | -29.1277
18.7645 | -35.1569 | -29.1347
18.7706 | -35.1574 | -29.1419
18.7767 | -35.1578 | -29.1491
18.7828 | -35.1584 | -29.156
18.7889 | -35.1588 | -29.1634
18.795 | -35.1593 | -29.1704
18.8011 | -35.1598 | -29.1779
18.8072 | -35.1603 | -29.1853
18.8133 | -35.1607 | -29.1926
18.8194 | -35.1612 | -29.2
18.8255 | -35.1617 | -29.2073
18.8316 | -35.1622 | -29.2149
18.8377 | -35.1626 | -29.2218
18.8439 | -35.1631 | -29.229
18.85 | -35.1635 | -29.2365
18.8561 | -35.164 | -29.2435
18.8622 | -35.1645 | -29.251
18.8683 | -35.1649 | -29.258
18.8744 | -35.1655 | -29.2652
18.8805 | -35.166 | -29.2727
18.8866 | -35.1664 | -29.2805
18.8927 | -35.167 | -29.2875
18.8988 | -35.1674 | -29.2944
18.9049 | -35.1679 | -29.3018
18.911 | -35.1684 | -29.3091
18.9171 | -35.1689 | -29.3164
18.9232 | -35.1694 | -29.3235
18.9293 | -35.1699 | -29.3306
18.9354 | -35.1703 | -29.3377
18.9415 | -35.1708 | -29.3446
18.9476 | -35.1713 | -29.3519
18.9537 | -35.1718 | -29.3587
18.9598 | -35.1724 | -29.3655
18.9659 | -35.1729 | -29.3725
18.972 | -35.1733 | -29.3797
18.9781 | -35.1738 | -29.3865
18.9843 | -35.1743 | -29.3939
18.9904 | -35.1747 | -29.4011
18.9965 | -35.1753 | -29.4081
19.0026 | -35.1758 | -29.4154
19.0087 | -35.1762 | -29.4229
19.0148 | -35.1768 | -29.4298
19.0209 | -35.1772 | -29.4374
19.027 | -35.1777 | -29.4449
19.0331 | -35.1782 | -29.4518
19.0392 | -35.1787 | -29.4589
19.0453 | -35.1792 | -29.4659
19.0514 | -35.1796 | -29.4728
19.0575 | -35.18 | -29.48
19.0636 | -35.1806 | -29.4866
19.0697 | -35.1811 | -29.4939
19.0758 | -35.1816 | -29.501
19.0819 | -35.1821 | -29.5085
19.088 | -35.1826 | -29.5155
19.0941 | -35.1831 | -29.5221
19.1002 | -35.1836 | -29.5291
19.1063 | -35.1841 | -29.5364
19.1124 | -35.1846 | -29.5432
19.1185 | -35.185 | -29.5504
19.1246 | -35.1854 | -29.5575
19.1308 | -35.1859 | -29.5644
19.1369 | -35.1864 | -29.5713
19.143 | -35.1869 | -29.5784
19.1491 | -35.1874 | -29.5852
19.1552 | -35.1879 | -29.5927
19.1613 | -35.1883 | -29.5997
19.1674 | -35.1888 | -29.6073
19.1735 | -35.1893 | -29.6142
19.1796 | -35.1898 | -29.6214
19.1857 | -35.1903 | -29.6288
19.1918 | -35.1907 | -29.6361
19.1979 | -35.1912 | -29.6435
19.204 | -35.1917 | -29.6501
19.2101 | -35.1922 | -29.6575
19.2162 | -35.1926 | -29.6648
19.2223 | -35.1931 | -29.6719
19.2284 | -35.1936 | -29.6796
19.2345 | -35.1941 | -29.6868
19.2406 | -35.1946 | -29.6939
19.2467 | -35.1951 | -29.7014
19.2528 | -35.1956 | -29.7087
19.2589 | -35.1961 | -29.7156
19.265 | -35.1966 | -29.7229
19.2712 | -35.1971 | -29.7302
19.2773 | -35.1976 | -29.737
19.2834 | -35.1981 | -29.7438
19.2895 | -35.1986 | -29.7509
19.2956 | -35.1991 | -29.7584
19.3017 | -35.1996 | -29.7656
19.3078 | -35.2001 | -29.7728
19.3139 | -35.2006 | -29.7798
19.32 | -35.2011 | -29.7869
19.3261 | -35.2016 | -29.7938
19.3322 | -35.2021 | -29.8008
19.3383 | -35.2026 | -29.8077
19.3444 | -35.2032 | -29.8141
19.3505 | -35.2036 | -29.8214
19.3566 | -35.204 | -29.8288
19.3627 | -35.2045 | -29.8361
19.3688 | -35.205 | -29.8436
19.3749 | -35.2055 | -29.8507
19.381 | -35.206 | -29.8578
19.3871 | -35.2065 | -29.8647
19.3932 | -35.207 | -29.8719
19.3993 | -35.2075 | -29.8791
19.4054 | -35.208 | -29.886
19.4115 | -35.2085 | -29.8934
19.4177 | -35.2089 | -29.9005
19.4238 | -35.2094 | -29.9074
19.4299 | -35.2099 | -29.9141
19.436 | -35.2104 | -29.9218
19.4421 | -35.2108 | -29.9287
19.4482 | -35.2113 | -29.9359
19.4543 | -35.2119 | -29.9424
19.4604 | -35.2123 | -29.9493
19.4665 | -35.2128 | -29.9564
19.4726 | -35.2133 | -29.9632
19.4787 | -35.2137 | -29.9703
19.4848 | -35.2142 | -29.9772
19.4909 | -35.2147 | -29.9838
19.497 | -35.2151 | -29.9908
19.5031 | -35.2156 | -29.9975
19.5092 | -35.216 | -30.0051
19.5153 | -35.2165 | -30.0125
19.5214 | -35.217 | -30.0194
19.5275 | -35.2175 | -30.0263
19.5336 | -35.2179 | -30.033
19.5397 | -35.2185 | -30.0398
19.5458 | -35.219 | -30.0469
19.5519 | -35.2194 | -30.054
19.5581 | -35.22 | -30.0602
19.5642 | -35.2205 | -30.0674
19.5703 | -35.221 | -30.074
19.5764 | -35.2215 | -30.0811
19.5825 | -35.2219 | -30.088
19.5886 | -35.2224 | -30.0948
19.5947 | -35.2229 | -30.1015
19.6008 | -35.2234 | -30.1085
19.6069 | -35.2239 | -30.1149
19.613 | -35.2244 | -30.1218
19.6191 | -35.2249 | -30.1285
19.6252 | -35.2254 | -30.135
19.6313 | -35.2258 | -30.1422
19.6374 | -35.2263 | -30.1494
19.6435 | -35.2268 | -30.156
19.6496 | -35.2273 | -30.1627
19.6557 | -35.2279 | -30.1696
19.6618 | -35.2285 | -30.1766
19.6679 | -35.229 | -30.1838
19.674 | -35.2294 | -30.1904
19.6801 | -35.2298 | -30.1977
19.6862 | -35.2303 | -30.2047
19.6923 | -35.2308 | -30.2117
19.6984 | -35.2313 | -30.2181
19.7046 | -35.2318 | -30.2248
19.7107 | -35.2323 | -30.2321
19.7168 | -35.2327 | -30.2393
19.7229 | -35.2332 | -30.247
19.729 | -35.2336 | -30.2546
19.7351 | -35.2343 | -30.2614
19.7412 | -35.2348 | -30.2685
19.7473 | -35.2353 | -30.2752
19.7534 | -35.2357 | -30.2824
19.7595 | -35.2362 | -30.2898
19.7656 | -35.2367 | -30.2973
19.7717 | -35.2372 | -30.3038
19.7778 | -35.2377 | -30.311
19.7839 | -35.2382 | -30.3176
19.79 | -35.2387 | -30.3248
19.7961 | -35.2391 | -30.3318
19.8022 | -35.2396 | -30.3387
19.8083 | -35.2401 | -30.3457
19.8144 | -35.2406 | -30.3528
19.8205 | -35.2411 | -30.3595
19.8266 | -35.2416 | -30.3666
19.8327 | -35.242 | -30.3736
19.8388 | -35.2425 | -30.3806
19.845 | -35.243 | -30.3873
19.8511 | -35.2436 | -30.3944
19.8572 | -35.244 | -30.4015
19.8633 | -35.2445 | -30.4084
19.8694 | -35.2451 | -30.4162
19.8755 | -35.2456 | -30.4237
19.8816 | -35.2461 | -30.4309
19.8877 | -35.2465 | -30.438
19.8938 | -35.2471 | -30.4448
19.8999 | -35.2475 | -30.4517
19.906 | -35.248 | -30.4588
19.9121 | -35.2486 | -30.4652
19.9182 | -35.2491 | -30.4714
19.9243 | -35.2497 | -30.4784
19.9304 | -35.2502 | -30.4855
19.9365 | -35.2507 | -30.4923
19.9426 | -35.2513 | -30.4985
19.9487 | -35.2518 | -30.5051
19.9548 | -35.2524 | -30.5127
19.9609 | -35.253 | -30.5203
19.967 | -35.2535 | -30.5268
19.9731 | -35.254 | -30.5337
19.9792 | -35.2545 | -30.5401
19.9853 | -35.2551 | -30.547
19.9915 | -35.2556 | -30.5543
19.9976 | -35.256 | -30.5611
20.0037 | -35.2566 | -30.5682
20.0098 | -35.257 | -30.5751
20.0159 | -35.2575 | -30.5822
20.022 | -35.258 | -30.5886
20.0281 | -35.2585 | -30.5958
20.0342 | -35.259 | -30.6024
20.0403 | -35.2594 | -30.6084
20.0464 | -35.2599 | -30.6161
20.0525 | -35.2603 | -30.6229
20.0586 | -35.2609 | -30.6301
20.0647 | -35.2613 | -30.638
20.0708 | -35.2618 | -30.6449
20.0769 | -35.2623 | -30.6521
20.083 | -35.2629 | -30.6589
20.0891 | -35.2633 | -30.6662
20.0952 | -35.2639 | -30.6731
20.1013 | -35.2642 | -30.6797
20.1074 | -35.2647 | -30.6866
20.1135 | -35.2652 | -30.6938
20.1196 | -35.2657 | -30.7015
20.1257 | -35.2662 | -30.7084
20.1319 | -35.2667 | -30.7151
20.138 | -35.2672 | -30.7217
20.1441 | -35.2677 | -30.7287
20.1502 | -35.2681 | -30.736
20.1563 | -35.2686 | -30.7421
20.1624 | -35.2692 | -30.749
20.1685 | -35.2697 | -30.7561
20.1746 | -35.2701 | -30.7629
20.1807 | -35.2707 | -30.7695
20.1868 | -35.2712 | -30.7766
20.1929 | -35.2716 | -30.7833
20.199 | -35.2721 | -30.7902
20.2051 | -35.2727 | -30.7968
20.2112 | -35.2732 | -30.8038
20.2173 | -35.2736 | -30.8105
20.2234 | -35.274 | -30.8174
20.2295 | -35.2746 | -30.8249
20.2356 | -35.2751 | -30.8323
20.2417 | -35.2756 | -30.8388
20.2478 | -35.276 | -30.8457
20.2539 | -35.2764 | -30.8519
20.26 | -35.277 | -30.8585
20.2661 | -35.2775 | -30.8654
20.2723 | -35.278 | -30.8721
20.2784 | -35.2785 | -30.8788
20.2845 | -35.2791 | -30.8859
20.2906 | -35.2796 | -30.8928
20.2967 | -35.2802 | -30.899
20.3028 | -35.2806 | -30.9057
20.3089 | -35.2812 | -30.9121
20.315 | -35.2816 | -30.9187
20.3211 | -35.2822 | -30.9261
20.3272 | -35.2827 | -30.9333
20.3333 | -35.2832 | -30.9403
20.3394 | -35.2838 | -30.9469
20.3455 | -35.2842 | -30.954
20.3516 | -35.2847 | -30.9609
20.3577 | -35.2851 | -30.9674
20.3638 | -35.2857 | -30.974
20.3699 | -35.2863 | -30.9794
20.376 | -35.2868 | -30.9863
20.3821 | -35.2873 | -30.9934
20.3882 | -35.2878 | -31.0009
20.3943 | -35.2883 | -31.0084
20.4004 | -35.2889 | -31.0148
20.4065 | -35.2894 | -31.0223
20.4126 | -35.2899 | -31.0287
20.4188 | -35.2903 | -31.0358
20.4249 | -35.2909 | -31.0429
20.431 | -35.2914 | -31.0501
20.4371 | -35.2918 | -31.056
20.4432 | -35.2923 | -31.0629
20.4493 | -35.2929 | -31.0695
20.4554 | -35.2934 | -31.076
20.4615 | -35.2939 | -31.0835
20.4676 | -35.2945 | -31.0904
20.4737 | -35.295 | -31.0973
20.4798 | -35.2954 | -31.1046
20.4859 | -35.2959 | -31.1108
20.492 | -35.2965 | -31.1183
20.4981 | -35.297 | -31.1244
20.5042 | -35.2974 | -31.1311
20.5103 | -35.2979 | -31.137
20.5164 | -35.2984 | -31.1437
20.5225 | -35.2989 | -31.1502
20.5286 | -35.2993 | -31.158
20.5347 | -35.2997 | -31.1648
20.5408 | -35.3002 | -31.1716
20.5469 | -35.3008 | -31.1787
20.553 | -35.3012 | -31.1852
20.5592 | -35.3017 | -31.1918
20.5653 | -35.3021 | -31.1982
20.5714 | -35.3026 | -31.2052
20.5775 | -35.3031 | -31.2118
20.5836 | -35.3036 | -31.2183
20.5897 | -35.3041 | -31.2252
20.5958 | -35.3046 | -31.2321
20.6019 | -35.3052 | -31.2386
20.608 | -35.3057 | -31.2453
20.6141 | -35.3062 | -31.2521
20.6202 | -35.3066 | -31.2586
20.6263 | -35.3072 | -31.2656
20.6324 | -35.3076 | -31.2718
20.6385 | -35.3081 | -31.2783
20.6446 | -35.3086 | -31.2848
20.6507 | -35.309 | -31.2918
20.6568 | -35.3095 | -31.2983
20.6629 | -35.31 | -31.3046
20.669 | -35.3104 | -31.3118
20.6751 | -35.3108 | -31.3188
20.6812 | -35.3112 | -31.3249
20.6873 | -35.3118 | -31.3312
20.6934 | -35.3125 | -31.3375
20.6995 | -35.313 | -31.3441
20.7057 | -35.3135 | -31.3511
20.7118 | -35.314 | -31.3584
20.7179 | -35.3146 | -31.3641
20.724 | -35.315 | -31.3705
20.7301 | -35.3154 | -31.3783
20.7362 | -35.316 | -31.3857
20.7423 | -35.3166 | -31.3919
20.7484 | -35.3171 | -31.3986
20.7545 | -35.3175 | -31.405
20.7606 | -35.318 | -31.4118
20.7667 | -35.3186 | -31.4179
20.7728 | -35.3192 | -31.4243
20.7789 | -35.3196 | -31.4306
20.785 | -35.32 | -31.4376
20.7911 | -35.3206 | -31.4434
20.7972 | -35.3211 | -31.45
20.8033 | -35.3216 | -31.457
20.8094 | -35.3222 | -31.4633
20.8155 | -35.3226 | -31.4705
20.8216 | -35.323 | -31.4769
20.8277 | -35.3235 | -31.4834
20.8338 | -35.3239 | -31.4901
20.8399 | -35.3245 | -31.4974
20.8461 | -35.325 | -31.5042
20.8522 | -35.3255 | -31.5107
20.8583 | -35.3258 | -31.5179
20.8644 | -35.3263 | -31.5245
20.8705 | -35.3268 | -31.5306
20.8766 | -35.3274 | -31.5379
20.8827 | -35.3279 | -31.5451
20.8888 | -35.3284 | -31.5527
20.8949 | -35.3289 | -31.5589
20.901 | -35.3294 | -31.5659
20.9071 | -35.3299 | -31.5721
20.9132 | -35.3304 | -31.5786
20.9193 | -35.3308 | -31.5847
20.9254 | -35.3314 | -31.5914
20.9315 | -35.332 | -31.5967
20.9376 | -35.3325 | -31.603
20.9437 | -35.333 | -31.6098
20.9498 | -35.3336 | -31.6157
20.9559 | -35.3341 | -31.6219
20.962 | -35.3346 | -31.6289
20.9681 | -35.3351 | -31.6356
20.9742 | -35.3355 | -31.6421
20.9803 | -35.336 | -31.649
20.9864 | -35.3365 | -31.6553
20.9926 | -35.3369 | -31.6624
20.9987 | -35.3375 | -31.6682
21.0048 | -35.3379 | -31.675
21.0109 | -35.3385 | -31.6817
21.017 | -35.3388 | -31.6879
21.0231 | -35.3393 | -31.6953
21.0292 | -35.3399 | -31.7018
21.0353 | -35.3404 | -31.7084
21.0414 | -35.3409 | -31.7147
21.0475 | -35.3415 | -31.7212
21.0536 | -35.342 | -31.728
21.0597 | -35.3425 | -31.7348
21.0658 | -35.343 | -31.7419
21.0719 | -35.3435 | -31.7488
21.078 | -35.344 | -31.7553
21.0841 | -35.3444 | -31.7615
21.0902 | -35.3447 | -31.7671
21.0963 | -35.3451 | -31.7733
21.1024 | -35.3456 | -31.779
21.1085 | -35.3463 | -31.7848
21.1146 | -35.3468 | -31.7909
21.1207 | -35.3474 | -31.7969
21.1268 | -35.3478 | -31.8038
21.133 | -35.3484 | -31.8101
21.1391 | -35.3489 | -31.8151
21.1452 | -35.3496 | -31.8214
21.1513 | -35.3501 | -31.8276
21.1574 | -35.3505 | -31.8337
21.1635 | -35.351 | -31.8403
21.1696 | -35.3513 | -31.8471
21.1757 | -35.3519 | -31.852
21.1818 | -35.3526 | -31.8587
21.1879 | -35.3532 | -31.8645
21.194 | -35.3538 | -31.8716
21.2001 | -35.3542 | -31.8781
21.2062 | -35.3546 | -31.885
21.2123 | -35.355 | -31.8935
21.2184 | -35.3553 | -31.9005
21.2245 | -35.3558 | -31.9074
21.2306 | -35.3562 | -31.9141
21.2367 | -35.3568 | -31.9215
21.2428 | -35.3572 | -31.9283
21.2489 | -35.3577 | -31.9358
21.255 | -35.3583 | -31.9431
21.2611 | -35.3588 | -31.9492
21.2672 | -35.3592 | -31.9555
21.2733 | -35.3597 | -31.9615
21.2795 | -35.3599 | -31.9685
21.2856 | -35.3603 | -31.9756
21.2917 | -35.3608 | -31.983
21.2978 | -35.3612 | -31.9899
21.3039 | -35.3616 | -31.9953
21.31 | -35.3621 | -32.001
21.3161 | -35.3625 | -32.0073
21.3222 | -35.3631 | -32.0133
21.3283 | -35.3635 | -32.0197
21.3344 | -35.364 | -32.0259
21.3405 | -35.3644 | -32.0329
21.3466 | -35.3648 | -32.0386
21.3527 | -35.3653 | -32.0451
21.3588 | -35.3659 | -32.0511
21.3649 | -35.3663 | -32.058
21.371 | -35.3668 | -32.0645
21.3771 | -35.3673 | -32.0702
21.3832 | -35.3677 | -32.0764
21.3893 | -35.3683 | -32.0822
21.3954 | -35.3688 | -32.0885
21.4015 | -35.3693 | -32.0956
21.4076 | -35.3698 | -32.1021
21.4137 | -35.3703 | -32.1087
21.4199 | -35.3707 | -32.1155
21.426 | -35.3711 | -32.122
21.4321 | -35.3717 | -32.1285
21.4382 | -35.3722 | -32.1354
21.4443 | -35.3726 | -32.1409
21.4504 | -35.3732 | -32.1481
21.4565 | -35.3736 | -32.1549
21.4626 | -35.3742 | -32.1622
21.4687 | -35.3747 | -32.1686
21.4748 | -35.3753 | -32.1755
21.4809 | -35.3757 | -32.1822
21.487 | -35.3762 | -32.1886
21.4931 | -35.3767 | -32.1945
21.4992 | -35.3772 | -32.2
21.5053 | -35.3778 | -32.207
21.5114 | -35.3783 | -32.213
21.5175 | -35.3787 | -32.2203
21.5236 | -35.3791 | -32.2268
21.5297 | -35.3796 | -32.2317
21.5358 | -35.3801 | -32.2382
21.5419 | -35.3805 | -32.2452
21.548 | -35.3809 | -32.2516
21.5541 | -35.3813 | -32.2583
21.5602 | -35.3818 | -32.2641
21.5664 | -35.3822 | -32.2698
21.5725 | -35.3826 | -32.2762
21.5786 | -35.3831 | -32.2827
21.5847 | -35.3836 | -32.2896
21.5908 | -35.3841 | -32.2968
21.5969 | -35.3844 | -32.3028
21.603 | -35.3849 | -32.3094
21.6091 | -35.3854 | -32.3162
21.6152 | -35.386 | -32.3225
21.6213 | -35.3863 | -32.3294
21.6274 | -35.3869 | -32.3358
21.6335 | -35.3874 | -32.342
21.6396 | -35.3878 | -32.3489
21.6457 | -35.3883 | -32.3552
21.6518 | -35.3887 | -32.3613
21.6579 | -35.3892 | -32.3679
21.664 | -35.3896 | -32.3744
21.6701 | -35.3902 | -32.3793
21.6762 | -35.3907 | -32.3858
21.6823 | -35.3911 | -32.3909
21.6884 | -35.3915 | -32.3974
21.6945 | -35.3919 | -32.404
21.7006 | -35.3925 | -32.4097
21.7068 | -35.3929 | -32.4153
21.7129 | -35.3934 | -32.4225
21.719 | -35.394 | -32.4281
21.7251 | -35.3946 | -32.4336
21.7312 | -35.3951 | -32.4394
21.7373 | -35.3957 | -32.4453
21.7434 | -35.3961 | -32.4506
21.7495 | -35.3966 | -32.4573
21.7556 | -35.3972 | -32.4635
21.7617 | -35.3977 | -32.4699
21.7678 | -35.3983 | -32.4765
21.7739 | -35.3987 | -32.4827
21.78 | -35.3994 | -32.489
21.7861 | -35.3999 | -32.4945
21.7922 | -35.4003 | -32.5001
21.7983 | -35.401 | -32.5059
21.8044 | -35.4016 | -32.5122
21.8105 | -35.4021 | -32.518
21.8166 | -35.4025 | -32.5244
21.8227 | -35.4029 | -32.5302
21.8288 | -35.4033 | -32.5364
21.8349 | -35.4039 | -32.5424
21.841 | -35.4044 | -32.5487
21.8471 | -35.405 | -32.5538
21.8533 | -35.4055 | -32.5599
21.8594 | -35.4059 | -32.5668
21.8655 | -35.4064 | -32.5741
21.8716 | -35.4069 | -32.5805
21.8777 | -35.4074 | -32.5868
21.8838 | -35.4078 | -32.5928
21.8899 | -35.4084 | -32.5985
21.896 | -35.4089 | -32.6049
21.9021 | -35.4093 | -32.611
21.9082 | -35.4097 | -32.6173
21.9143 | -35.4101 | -32.6231
21.9204 | -35.4105 | -32.6298
21.9265 | -35.411 | -32.6357
21.9326 | -35.4114 | -32.6419
21.9387 | -35.4118 | -32.6481
21.9448 | -35.4123 | -32.6543
21.9509 | -35.4129 | -32.6603
21.957 | -35.4134 | -32.6665
21.9631 | -35.4137 | -32.6728
21.9692 | -35.4143 | -32.6797
21.9753 | -35.4147 | -32.6859
21.9814 | -35.4151 | -32.6936
21.9875 | -35.4155 | -32.6999
21.9937 | -35.416 | -32.7055
21.9998 | -35.4163 | -32.7111
22.0059 | -35.4167 | -32.7181
22.012 | -35.4171 | -32.7238
22.0181 | -35.4175 | -32.7301
22.0242 | -35.418 | -32.737
22.0303 | -35.4183 | -32.7439
22.0364 | -35.4187 | -32.7487
22.0425 | -35.419 | -32.7544
22.0486 | -35.4196 | -32.7605
22.0547 | -35.4198 | -32.7679
22.0608 | -35.4202 | -32.7746
22.0669 | -35.4208 | -32.7813
22.073 | -35.4212 | -32.7875
22.0791 | -35.4217 | -32.7948
22.0852 | -35.4221 | -32.8014
22.0913 | -35.4226 | -32.8081
22.0974 | -35.4231 | -32.8132
22.1035 | -35.4234 | -32.8193
22.1096 | -35.4238 | -32.8259
22.1157 | -35.4242 | -32.8319
22.1218 | -35.4247 | -32.8385
22.1279 | -35.4251 | -32.8444
22.134 | -35.4255 | -32.8506
22.1402 | -35.4261 | -32.8558
22.1463 | -35.4266 | -32.8619
22.1524 | -35.4269 | -32.8675
22.1585 | -35.4275 | -32.8736
22.1646 | -35.4282 | -32.8796
22.1707 | -35.4287 | -32.8868
22.1768 | -35.4291 | -32.8924
22.1829 | -35.4296 | -32.8979
22.189 | -35.4298 | -32.9043
22.1951 | -35.4304 | -32.9107
22.2012 | -35.4308 | -32.9168
22.2073 | -35.4313 | -32.9231
22.2134 | -35.4318 | -32.9289
22.2195 | -35.4324 | -32.9355
22.2256 | -35.4328 | -32.9418
22.2317 | -35.4331 | -32.949
22.2378 | -35.4336 | -32.9553
22.2439 | -35.4341 | -32.9615
22.25 | -35.4347 | -32.9678
22.2561 | -35.4353 | -32.9734
22.2622 | -35.4358 | -32.9796
22.2683 | -35.4363 | -32.9851
22.2744 | -35.4368 | -32.9913
22.2806 | -35.4373 | -32.9977
22.2867 | -35.4378 | -33.0045
22.2928 | -35.4383 | -33.0106
22.2989 | -35.4388 | -33.0165
22.305 | -35.4391 | -33.0224
22.3111 | -35.4397 | -33.0293
22.3172 | -35.4402 | -33.0354
22.3233 | -35.4407 | -33.041
22.3294 | -35.4411 | -33.0478
22.3355 | -35.4418 | -33.0539
22.3416 | -35.4422 | -33.0603
22.3477 | -35.4427 | -33.0666
22.3538 | -35.4431 | -33.073
22.3599 | -35.4435 | -33.0791
22.366 | -35.4439 | -33.0854
22.3721 | -35.4442 | -33.0914
22.3782 | -35.4446 | -33.0975
22.3843 | -35.4451 | -33.1035
22.3904 | -35.4457 | -33.1106
22.3965 | -35.4461 | -33.1169
22.4026 | -35.4467 | -33.1233
22.4087 | -35.447 | -33.1301
22.4148 | -35.4476 | -33.1371
22.4209 | -35.448 | -33.1441
22.4271 | -35.4483 | -33.1506
22.4332 | -35.4488 | -33.1568
22.4393 | -35.4492 | -33.1634
22.4454 | -35.4497 | -33.1684
22.4515 | -35.4501 | -33.1741
22.4576 | -35.4505 | -33.18
22.4637 | -35.4509 | -33.1859
22.4698 | -35.4515 | -33.1923
22.4759 | -35.4518 | -33.1976
22.482 | -35.4522 | -33.2036
22.4881 | -35.4527 | -33.2107
22.4942 | -35.4531 | -33.2163
22.5003 | -35.4535 | -33.2218
22.5064 | -35.454 | -33.2279
22.5125 | -35.4544 | -33.2333
22.5186 | -35.455 | -33.2385
22.5247 | -35.4555 | -33.2453
22.5308 | -35.4559 | -33.2507
22.5369 | -35.4563 | -33.2567
22.543 | -35.4566 | -33.2634
22.5491 | -35.4569 | -33.2685
22.5552 | -35.4573 | -33.2741
22.5613 | -35.4577 | -33.2805
22.5675 | -35.4581 | -33.2876
22.5736 | -35.4587 | -33.2942
22.5797 | -35.4593 | -33.2999
22.5858 | -35.4596 | -33.3065
22.5919 | -35.4601 | -33.3121
22.598 | -35.4605 | -33.3173
22.6041 | -35.4609 | -33.3236
22.6102 | -35.4615 | -33.3299
22.6163 | -35.462 | -33.3359
22.6224 | -35.4625 | -33.3418
22.6285 | -35.463 | -33.3477
22.6346 | -35.4634 | -33.3547
22.6407 | -35.4639 | -33.3606
22.6468 | -35.4643 | -33.3679
22.6529 | -35.4647 | -33.3731
22.659 | -35.4652 | -33.3792
22.6651 | -35.4656 | -33.3851
22.6712 | -35.466 | -33.3917
22.6773 | -35.4665 | -33.3967
22.6834 | -35.467 | -33.4041
22.6895 | -35.4674 | -33.4102
22.6956 | -35.468 | -33.4155
22.7017 | -35.4685 | -33.4213
22.7079 | -35.4688 | -33.4273
22.714 | -35.4692 | -33.4334
22.7201 | -35.4697 | -33.4392
22.7262 | -35.4701 | -33.4453
22.7323 | -35.4706 | -33.4511
22.7384 | -35.471 | -33.4568
22.7445 | -35.4713 | -33.4626
22.7506 | -35.4717 | -33.4684
22.7567 | -35.4723 | -33.4738
22.7628 | -35.4726 | -33.4795
22.7689 | -35.473 | -33.4856
22.775 | -35.4733 | -33.4916
22.7811 | -35.4737 | -33.4971
22.7872 | -35.4743 | -33.5037
22.7933 | -35.4747 | -33.5097
22.7994 | -35.4751 | -33.5162
22.8055 | -35.4755 | -33.5218
22.8116 | -35.4759 | -33.5281
22.8177 | -35.4764 | -33.5341
22.8238 | -35.477 | -33.5404
22.8299 | -35.4774 | -33.5469
22.836 | -35.4779 | -33.5537
22.8421 | -35.4784 | -33.5594
22.8482 | -35.4788 | -33.5653
22.8544 | -35.4792 | -33.5714
22.8605 | -35.4798 | -33.5778
22.8666 | -35.4801 | -33.5841
22.8727 | -35.4806 | -33.5909
22.8788 | -35.4812 | -33.5965
22.8849 | -35.4815 | -33.6019
22.891 | -35.4821 | -33.6074
22.8971 | -35.4826 | -33.6141
22.9032 | -35.4831 | -33.6204
22.9093 | -35.4835 | -33.6261
22.9154 | -35.484 | -33.631
22.9215 | -35.4844 | -33.6377
22.9276 | -35.4847 | -33.6428
22.9337 | -35.4848 | -33.6476
22.9398 | -35.4853 | -33.6526
22.9459 | -35.4856 | -33.6577
22.952 | -35.486 | -33.6624
22.9581 | -35.4864 | -33.6683
22.9642 | -35.4868 | -33.6746
22.9703 | -35.4872 | -33.6806
22.9764 | -35.4875 | -33.6864
22.9825 | -35.4879 | -33.6922
22.9886 | -35.4883 | -33.6986
22.9948 | -35.4886 | -33.7039
23.0009 | -35.489 | -33.7101
23.007 | -35.4895 | -33.7145
23.0131 | -35.49 | -33.7199
23.0192 | -35.4905 | -33.7253
23.0253 | -35.4909 | -33.7306
23.0314 | -35.4913 | -33.7371
23.0375 | -35.4917 | -33.7433
23.0436 | -35.4921 | -33.7497
23.0497 | -35.4925 | -33.7556
23.0558 | -35.4929 | -33.7616
23.0619 | -35.4935 | -33.7671
23.068 | -35.494 | -33.773
23.0741 | -35.4944 | -33.7788
23.0802 | -35.4949 | -33.7843
23.0863 | -35.4953 | -33.7901
23.0924 | -35.4958 | -33.7961
23.0985 | -35.4962 | -33.8026
23.1046 | -35.4968 | -33.8083
23.1107 | -35.4973 | -33.8146
23.1168 | -35.4979 | -33.8201
23.1229 | -35.4983 | -33.8259
23.129 | -35.4987 | -33.8317
23.1351 | -35.4992 | -33.8375
23.1413 | -35.4996 | -33.8427
23.1474 | -35.5001 | -33.85
23.1535 | -35.5005 | -33.8559
23.1596 | -35.5011 | -33.8613
23.1657 | -35.5015 | -33.8668
23.1718 | -35.5019 | -33.8728
23.1779 | -35.5023 | -33.8784
23.184 | -35.5028 | -33.8838
23.1901 | -35.5033 | -33.8902
23.1962 | -35.5036 | -33.8959
23.2023 | -35.504 | -33.9008
23.2084 | -35.5045 | -33.9067
23.2145 | -35.5048 | -33.9119
23.2206 | -35.5053 | -33.9173
23.2267 | -35.5057 | -33.9233
23.2328 | -35.5061 | -33.9288
23.2389 | -35.5066 | -33.9358
23.245 | -35.507 | -33.9413
23.2511 | -35.5075 | -33.9481
23.2572 | -35.508 | -33.9533
23.2633 | -35.5083 | -33.9594
23.2694 | -35.5087 | -33.9653
23.2755 | -35.5091 | -33.9715
23.2817 | -35.5095 | -33.9776
23.2878 | -35.5099 | -33.9824
23.2939 | -35.5103 | -33.988
23.3 | -35.5108 | -33.9941
23.3061 | -35.5111 | -33.9995
23.3122 | -35.5116 | -34.0059
23.3183 | -35.512 | -34.0121
23.3244 | -35.5124 | -34.0187
23.3305 | -35.5128 | -34.0248
23.3366 | -35.5132 | -34.0306
23.3427 | -35.5137 | -34.0361
23.3488 | -35.5141 | -34.0415
23.3549 | -35.5146 | -34.0476
23.361 | -35.515 | -34.053
23.3671 | -35.5155 | -34.0589
23.3732 | -35.5159 | -34.0643
23.3793 | -35.5164 | -34.0705
23.3854 | -35.5168 | -34.0759
23.3915 | -35.5173 | -34.0821
23.3976 | -35.5177 | -34.0885
23.4037 | -35.5183 | -34.0948
23.4098 | -35.5188 | -34.1004
23.4159 | -35.5192 | -34.1061
23.422 | -35.5196 | -34.1106
23.4282 | -35.52 | -34.1166
23.4343 | -35.5204 | -34.122
23.4404 | -35.5209 | -34.1282
23.4465 | -35.5212 | -34.1335
23.4526 | -35.5216 | -34.1401
23.4587 | -35.5218 | -34.1467
23.4648 | -35.5222 | -34.1525
23.4709 | -35.5226 | -34.1578
23.477 | -35.523 | -34.1641
23.4831 | -35.5233 | -34.1695
23.4892 | -35.5237 | -34.1755
23.4953 | -35.5241 | -34.181
23.5014 | -35.5247 | -34.1865
23.5075 | -35.5251 | -34.1924
23.5136 | -35.5256 | -34.1987
23.5197 | -35.526 | -34.2057
23.5258 | -35.5264 | -34.2103
23.5319 | -35.5267 | -34.2158
23.538 | -35.5269 | -34.2213
23.5441 | -35.5272 | -34.227
23.5502 | -35.5277 | -34.2321
23.5563 | -35.5282 | -34.2379
23.5624 | -35.5286 | -34.2435
23.5686 | -35.5293 | -34.2505
23.5747 | -35.5296 | -34.256
23.5808 | -35.5299 | -34.2615
23.5869 | -35.5305 | -34.2678
23.593 | -35.5309 | -34.2734
23.5991 | -35.5315 | -34.2793
23.6052 | -35.5321 | -34.2843
23.6113 | -35.5326 | -34.2902
23.6174 | -35.5331 | -34.2955
23.6235 | -35.5335 | -34.3009
23.6296 | -35.5339 | -34.3064
23.6357 | -35.5344 | -34.3124
23.6418 | -35.5348 | -34.3187
23.6479 | -35.5352 | -34.3244
23.654 | -35.5354 | -34.3309
23.6601 | -35.536 | -34.3365
23.6662 | -35.5366 | -34.3417
23.6723 | -35.537 | -34.3467
23.6784 | -35.5376 | -34.3535
23.6845 | -35.538 | -34.3598
23.6906 | -35.5383 | -34.3652
23.6967 | -35.5387 | -34.3706
23.7028 | -35.5392 | -34.3771
23.7089 | -35.5397 | -34.3839
23.7151 | -35.5402 | -34.3895
23.7212 | -35.5407 | -34.3959
23.7273 | -35.5409 | -34.4009
23.7334 | -35.5413 | -34.4073
23.7395 | -35.5418 | -34.4121
23.7456 | -35.5425 | -34.417
23.7517 | -35.5429 | -34.4224
23.7578 | -35.5432 | -34.428
23.7639 | -35.5436 | -34.4345
23.77 | -35.544 | -34.4403
23.7761 | -35.5446 | -34.4464
23.7822 | -35.5451 | -34.4524
23.7883 | -35.5455 | -34.458
23.7944 | -35.5459 | -34.4639
23.8005 | -35.5463 | -34.4711
23.8066 | -35.5467 | -34.4761
23.8127 | -35.5471 | -34.4811
23.8188 | -35.5476 | -34.4864
23.8249 | -35.5481 | -34.4914
23.831 | -35.5485 | -34.4975
23.8371 | -35.5489 | -34.5028
23.8432 | -35.5494 | -34.508
23.8493 | -35.55 | -34.5129
23.8555 | -35.5504 | -34.5192
23.8616 | -35.5507 | -34.526
23.8677 | -35.5512 | -34.5306
23.8738 | -35.5515 | -34.5365
23.8799 | -35.5521 | -34.5423
23.886 | -35.5527 | -34.5484
23.8921 | -35.553 | -34.5557
23.8982 | -35.5534 | -34.5607
23.9043 | -35.5537 | -34.5666
23.9104 | -35.5542 | -34.5713
23.9165 | -35.5545 | -34.576
23.9226 | -35.5548 | -34.5814
23.9287 | -35.5552 | -34.5873
23.9348 | -35.5555 | -34.5918
23.9409 | -35.5558 | -34.5988
23.947 | -35.556 | -34.6031
23.9531 | -35.5566 | -34.61
23.9592 | -35.557 | -34.6164
23.9653 | -35.5575 | -34.6221
23.9714 | -35.5581 | -34.6277
23.9775 | -35.5585 | -34.6341
23.9836 | -35.559 | -34.6404
23.9897 | -35.5593 | -34.6464
23.9958 | -35.5598 | -34.6528
24.002 | -35.5605 | -34.6587
24.0081 | -35.5609 | -34.6632
24.0142 | -35.5613 | -34.6678
24.0203 | -35.5618 | -34.6738
24.0264 | -35.5618 | -34.6798
24.0325 | -35.5623 | -34.6859
24.0386 | -35.5627 | -34.6916
24.0447 | -35.5631 | -34.6966
24.0508 | -35.5635 | -34.7023
24.0569 | -35.564 | -34.707
24.063 | -35.5644 | -34.7116
24.0691 | -35.5649 | -34.7167
24.0752 | -35.5653 | -34.7221
24.0813 | -35.5658 | -34.7277
24.0874 | -35.5663 | -34.7333
24.0935 | -35.5667 | -34.739
24.0996 | -35.5673 | -34.7439
24.1057 | -35.5677 | -34.7492
24.1118 | -35.5681 | -34.7552
24.1179 | -35.5686 | -34.7607
24.124 | -35.569 | -34.7667
24.1301 | -35.5694 | -34.7722
24.1362 | -35.5697 | -34.7783
24.1424 | -35.5702 | -34.783
24.1485 | -35.5705 | -34.7883
24.1546 | -35.5709 | -34.7935
24.1607 | -35.5713 | -34.7998
24.1668 | -35.5719 | -34.8058
24.1729 | -35.5721 | -34.8116
24.179 | -35.5724 | -34.8167
24.1851 | -35.5729 | -34.8214
24.1912 | -35.5731 | -34.8273
24.1973 | -35.5736 | -34.833
24.2034 | -35.5744 | -34.837
24.2095 | -35.5749 | -34.8426
24.2156 | -35.5753 | -34.8484
24.2217 | -35.5758 | -34.8549
24.2278 | -35.5761 | -34.8605
24.2339 | -35.5765 | -34.8661
24.24 | -35.5769 | -34.8722
24.2461 | -35.5773 | -34.8776
24.2522 | -35.5776 | -34.8825
24.2583 | -35.578 | -34.8885
24.2644 | -35.5785 | -34.8956
24.2705 | -35.5788 | -34.9004
24.2766 | -35.5793 | -34.9069
24.2827 | -35.5798 | -34.9126
24.2889 | -35.5803 | -34.9184
24.295 | -35.5809 | -34.9238
24.3011 | -35.5813 | -34.93
24.3072 | -35.5818 | -34.9339
24.3133 | -35.5821 | -34.9393
24.3194 | -35.5825 | -34.9446
24.3255 | -35.5829 | -34.9509
24.3316 | -35.5834 | -34.9576
24.3377 | -35.5838 | -34.9633
24.3438 | -35.5841 | -34.9692
24.3499 | -35.5844 | -34.9748
24.356 | -35.5849 | -34.9794
24.3621 | -35.5853 | -34.9853
24.3682 | -35.5858 | -34.9911
24.3743 | -35.5863 | -34.9975
24.3804 | -35.5866 | -35.0038
24.3865 | -35.5869 | -35.0094
24.3926 | -35.5875 | -35.0158
24.3987 | -35.5879 | -35.0215
24.4048 | -35.5885 | -35.0263
24.4109 | -35.5891 | -35.0315
24.417 | -35.5894 | -35.0372
24.4231 | -35.5899 | -35.0424
24.4293 | -35.5903 | -35.0467
24.4354 | -35.5908 | -35.0528
24.4415 | -35.5913 | -35.0589
24.4476 | -35.5919 | -35.0642
24.4537 | -35.5924 | -35.0693
24.4598 | -35.5928 | -35.0739
24.4659 | -35.5932 | -35.0798
24.472 | -35.5936 | -35.0844
24.4781 | -35.594 | -35.0892
24.4842 | -35.5942 | -35.0941
24.4903 | -35.5946 | -35.0997
24.4964 | -35.5948 | -35.105
24.5025 | -35.5953 | -35.1126
24.5086 | -35.5957 | -35.1192
24.5147 | -35.5961 | -35.1245
24.5208 | -35.5966 | -35.1296
24.5269 | -35.5971 | -35.1353
24.533 | -35.5975 | -35.1423
24.5391 | -35.598 | -35.1482
24.5452 | -35.5984 | -35.1542
24.5513 | -35.5989 | -35.1605
24.5574 | -35.5994 | -35.1661
24.5635 | -35.5999 | -35.1719
24.5696 | -35.6004 | -35.1786
24.5758 | -35.601 | -35.1855
24.5819 | -35.6014 | -35.1917
24.588 | -35.6019 | -35.1971
24.5941 | -35.6024 | -35.2032
24.6002 | -35.603 | -35.2092
24.6063 | -35.6034 | -35.2141
24.6124 | -35.6038 | -35.2194
24.6185 | -35.6041 | -35.2252
24.6246 | -35.6043 | -35.2305
24.6307 | -35.6045 | -35.2382
24.6368 | -35.605 | -35.246
24.6429 | -35.6054 | -35.2519
24.649 | -35.6058 | -35.2573
24.6551 | -35.6062 | -35.263
24.6612 | -35.6068 | -35.2682
24.6673 | -35.6071 | -35.2745
24.6734 | -35.6074 | -35.2802
24.6795 | -35.6078 | -35.2865
24.6856 | -35.6082 | -35.2924
24.6917 | -35.6087 | -35.2971
24.6978 | -35.609 | -35.303
24.7039 | -35.6094 | -35.3096
24.71 | -35.61 | -35.315
24.7162 | -35.6104 | -35.3216
24.7223 | -35.6107 | -35.3277
24.7284 | -35.6112 | -35.3333
24.7345 | -35.6119 | -35.3381
24.7406 | -35.6123 | -35.3436
24.7467 | -35.6127 | -35.3497
24.7528 | -35.6131 | -35.356
24.7589 | -35.6134 | -35.3614
24.765 | -35.6139 | -35.3655
24.7711 | -35.6144 | -35.3717
24.7772 | -35.6147 | -35.3778
24.7833 | -35.6152 | -35.3832
24.7894 | -35.6157 | -35.3888
24.7955 | -35.6161 | -35.3944
24.8016 | -35.6168 | -35.4002
24.8077 | -35.6172 | -35.4065
24.8138 | -35.6179 | -35.4127
24.8199 | -35.6182 | -35.4197
24.826 | -35.6186 | -35.4244
24.8321 | -35.6191 | -35.4298
24.8382 | -35.6197 | -35.4353
24.8443 | -35.6201 | -35.4399
24.8504 | -35.6203 | -35.4447
24.8565 | -35.6208 | -35.4509
24.8627 | -35.6211 | -35.4563
24.8688 | -35.6217 | -35.4605
24.8749 | -35.6222 | -35.4666
24.881 | -35.6226 | -35.473
24.8871 | -35.6231 | -35.4785
24.8932 | -35.6236 | -35.484
24.8993 | -35.6238 | -35.4885
24.9054 | -35.6243 | -35.4945
24.9115 | -35.6246 | -35.5006
24.9176 | -35.625 | -35.5064
24.9237 | -35.6255 | -35.5131
24.9298 | -35.6258 | -35.5191
24.9359 | -35.6264 | -35.5248
24.942 | -35.6272 | -35.5308
24.9481 | -35.6276 | -35.5368
24.9542 | -35.628 | -35.5421
24.9603 | -35.6285 | -35.547
24.9664 | -35.6289 | -35.551
24.9725 | -35.6294 | -35.5571
24.9786 | -35.6298 | -35.5624
24.9847 | -35.63 | -35.5668
24.9908 | -35.6303 | -35.5726
24.9969 | -35.6308 | -35.5795
25.0031 | -35.6312 | -35.5835
25.0092 | -35.6318 | -35.5881
25.0153 | -35.6322 | -35.592
25.0214 | -35.6328 | -35.5975
25.0275 | -35.6332 | -35.6034
25.0336 | -35.6337 | -35.6098
25.0397 | -35.6342 | -35.6158
25.0458 | -35.6343 | -35.6208
25.0519 | -35.6348 | -35.6261
25.058 | -35.6352 | -35.6299
25.0641 | -35.6356 | -35.6349
25.0702 | -35.6359 | -35.6402
25.0763 | -35.6362 | -35.6468
25.0824 | -35.6364 | -35.6497
25.0885 | -35.6368 | -35.6548
25.0946 | -35.6374 | -35.6614
25.1007 | -35.6377 | -35.6669
25.1068 | -35.6382 | -35.6721
25.1129 | -35.6385 | -35.6775
25.119 | -35.6387 | -35.6829
25.1251 | -35.639 | -35.6891
25.1312 | -35.6395 | -35.6956
25.1373 | -35.6402 | -35.7013
25.1435 | -35.6406 | -35.7054
25.1496 | -35.6409 | -35.7103
25.1557 | -35.6413 | -35.7146
25.1618 | -35.6417 | -35.7212
25.1679 | -35.642 | -35.725
25.174 | -35.6423 | -35.731
25.1801 | -35.6429 | -35.7377
25.1862 | -35.6434 | -35.743
25.1923 | -35.6439 | -35.7474
25.1984 | -35.6443 | -35.7528
25.2045 | -35.6448 | -35.7588
25.2106 | -35.6454 | -35.7644
25.2167 | -35.6456 | -35.7702
25.2228 | -35.6461 | -35.7771
25.2289 | -35.6464 | -35.7818
25.235 | -35.6467 | -35.7883
25.2411 | -35.6472 | -35.7929
25.2472 | -35.6476 | -35.7985
25.2533 | -35.648 | -35.8036
25.2594 | -35.6488 | -35.8081
25.2655 | -35.6494 | -35.8134
25.2716 | -35.6498 | -35.8183
25.2777 | -35.6503 | -35.8225
25.2838 | -35.6506 | -35.8284
25.29 | -35.6514 | -35.8337
25.2961 | -35.6521 | -35.8381
25.3022 | -35.6526 | -35.844
25.3083 | -35.653 | -35.8514
25.3144 | -35.6532 | -35.8564
25.3205 | -35.6538 | -35.8628
25.3266 | -35.6541 | -35.8688
25.3327 | -35.6544 | -35.8742
25.3388 | -35.6548 | -35.8798
25.3449 | -35.655 | -35.8842
25.351 | -35.6551 | -35.889
25.3571 | -35.6554 | -35.8954
25.3632 | -35.6557 | -35.8999
25.3693 | -35.6564 | -35.9063
25.3754 | -35.6569 | -35.9123
25.3815 | -35.6573 | -35.9189
25.3876 | -35.6576 | -35.9247
25.3937 | -35.6581 | -35.9295
25.3998 | -35.6586 | -35.9347
25.4059 | -35.6589 | -35.9403
25.412 | -35.6594 | -35.9459
25.4181 | -35.6597 | -35.9513
25.4242 | -35.6602 | -35.9561
25.4304 | -35.6608 | -35.9613
25.4365 | -35.6611 | -35.9661
25.4426 | -35.6617 | -35.9717
25.4487 | -35.6619 | -35.9769
25.4548 | -35.6622 | -35.9838
25.4609 | -35.6626 | -35.9891
25.467 | -35.6631 | -35.9957
25.4731 | -35.6635 | -36.0012
25.4792 | -35.6638 | -36.0067
25.4853 | -35.6643 | -36.0116
25.4914 | -35.6649 | -36.0168
25.4975 | -35.6653 | -36.022
25.5036 | -35.666 | -36.027
25.5097 | -35.6665 | -36.0334
25.5158 | -35.6671 | -36.0403
25.5219 | -35.6677 | -36.0465
25.528 | -35.6681 | -36.0516
25.5341 | -35.6686 | -36.0578
25.5402 | -35.6691 | -36.0641
25.5463 | -35.6695 | -36.0702
25.5524 | -35.67 | -36.0753
25.5585 | -35.6706 | -36.0814
25.5646 | -35.6709 | -36.0883
25.5707 | -35.6711 | -36.0947
25.5769 | -35.6718 | -36.0993
25.583 | -35.6724 | -36.1055
25.5891 | -35.6729 | -36.1111
25.5952 | -35.6733 | -36.1161
25.6013 | -35.6737 | -36.1218
25.6074 | -35.6743 | -36.127
25.6135 | -35.6747 | -36.1335
25.6196 | -35.6752 | -36.1382
25.6257 | -35.6757 | -36.1438
25.6318 | -35.6761 | -36.1485
25.6379 | -35.6766 | -36.1538
25.644 | -35.677 | -36.1582
25.6501 | -35.6773 | -36.1639
25.6562 | -35.6776 | -36.17
25.6623 | -35.6779 | -36.1748
25.6684 | -35.6784 | -36.18
25.6745 | -35.6787 | -36.1861
25.6806 | -35.679 | -36.1922
25.6867 | -35.6794 | -36.1982
25.6928 | -35.6799 | -36.2026
25.6989 | -35.6803 | -36.2084
25.705 | -35.6809 | -36.2142
25.7111 | -35.6812 | -36.2203
25.7173 | -35.6817 | -36.2256
25.7234 | -35.6823 | -36.2305
25.7295 | -35.6826 | -36.236
25.7356 | -35.6831 | -36.2419
25.7417 | -35.6837 | -36.2466
25.7478 | -35.6842 | -36.2519
25.7539 | -35.6846 | -36.2581
25.76 | -35.685 | -36.2633
25.7661 | -35.6855 | -36.2679
25.7722 | -35.686 | -36.2731
25.7783 | -35.6864 | -36.2786
25.7844 | -35.6868 | -36.2841
25.7905 | -35.6873 | -36.2895
25.7966 | -35.6878 | -36.2945
25.8027 | -35.6883 | -36.3
25.8088 | -35.6888 | -36.3058
25.8149 | -35.6891 | -36.3114
25.821 | -35.6897 | -36.3164
25.8271 | -35.6902 | -36.3231
25.8332 | -35.6906 | -36.3282
25.8393 | -35.6909 | -36.3345
25.8454 | -35.6914 | -36.3393
25.8515 | -35.6919 | -36.3454
25.8576 | -35.6924 | -36.3503
25.8638 | -35.6929 | -36.3554
25.8699 | -35.6934 | -36.3609
25.876 | -35.6937 | -36.3661
25.8821 | -35.694 | -36.3712
25.8882 | -35.6944 | -36.3758
25.8943 | -35.6949 | -36.3817
25.9004 | -35.6954 | -36.389
25.9065 | -35.6957 | -36.3957
25.9126 | -35.6962 | -36.4015
25.9187 | -35.6966 | -36.4061
25.9248 | -35.6971 | -36.4113
25.9309 | -35.6976 | -36.4165
25.937 | -35.6981 | -36.4224
25.9431 | -35.6986 | -36.4279
25.9492 | -35.699 | -36.434
25.9553 | -35.6994 | -36.4394
25.9614 | -35.6997 | -36.4446
25.9675 | -35.7 | -36.4497
25.9736 | -35.7004 | -36.4556
25.9797 | -35.7008 | -36.4611
25.9858 | -35.701 | -36.4659
25.9919 | -35.7015 | -36.4716
25.998 | -35.702 | -36.4766
26.0042 | -35.7024 | -36.483
26.0103 | -35.703 | -36.4877
26.0164 | -35.7035 | -36.493
26.0225 | -35.7038 | -36.4981
26.0286 | -35.7042 | -36.5033
26.0347 | -35.7046 | -36.508
26.0408 | -35.705 | -36.514
26.0469 | -35.7055 | -36.5196
26.053 | -35.706 | -36.5249
26.0591 | -35.7066 | -36.5315
26.0652 | -35.7072 | -36.5355
26.0713 | -35.7075 | -36.5405
26.0774 | -35.7081 | -36.5461
26.0835 | -35.7085 | -36.5521
26.0896 | -35.7088 | -36.5569
26.0957 | -35.7093 | -36.5621
26.1018 | -35.7098 | -36.568
26.1079 | -35.7102 | -36.5729
26.114 | -35.7108 | -36.5791
26.1201 | -35.7111 | -36.5847
26.1262 | -35.7115 | -36.591
26.1323 | -35.7119 | -36.5969
26.1384 | -35.7125 | -36.6031
26.1445 | -35.7128 | -36.6093
26.1507 | -35.7133 | -36.615
26.1568 | -35.7138 | -36.6205
26.1629 | -35.7143 | -36.6256
26.169 | -35.7147 | -36.6319
26.1751 | -35.715 | -36.6368
26.1812 | -35.7154 | -36.642
26.1873 | -35.7159 | -36.6477
26.1934 | -35.7165 | -36.653
26.1995 | -35.717 | -36.6582
26.2056 | -35.7175 | -36.6638
26.2117 | -35.718 | -36.6702
26.2178 | -35.7184 | -36.6758
26.2239 | -35.7189 | -36.6808
26.23 | -35.7194 | -36.6859
26.2361 | -35.7199 | -36.6914
26.2422 | -35.7203 | -36.6963
26.2483 | -35.7209 | -36.7022
26.2544 | -35.7213 | -36.708
26.2605 | -35.7219 | -36.7141
26.2666 | -35.7222 | -36.7202
26.2727 | -35.7226 | -36.7264
26.2788 | -35.723 | -36.7313
26.2849 | -35.7234 | -36.7361
26.2911 | -35.7238 | -36.7418
26.2972 | -35.7242 | -36.7478
26.3033 | -35.7247 | -36.7534
26.3094 | -35.7251 | -36.7592
26.3155 | -35.7255 | -36.7639
26.3216 | -35.726 | -36.7695
26.3277 | -35.7264 | -36.7751
26.3338 | -35.7268 | -36.7806
26.3399 | -35.7273 | -36.786
26.346 | -35.7278 | -36.7914
26.3521 | -35.7282 | -36.7977
26.3582 | -35.7287 | -36.802
26.3643 | -35.7292 | -36.8084
26.3704 | -35.7297 | -36.8145
26.3765 | -35.7301 | -36.8204
26.3826 | -35.7306 | -36.8254
26.3887 | -35.7309 | -36.831
26.3948 | -35.7313 | -36.8372
26.4009 | -35.7319 | -36.843
26.407 | -35.7324 | -36.8489
26.4131 | -35.7327 | -36.8542
26.4192 | -35.7331 | -36.8593
26.4253 | -35.7337 | -36.864
26.4314 | -35.7341 | -36.8701
26.4376 | -35.7346 | -36.8757
26.4437 | -35.735 | -36.8807
26.4498 | -35.7355 | -36.8864
26.4559 | -35.736 | -36.891
26.462 | -35.7365 | -36.8963
26.4681 | -35.7369 | -36.9014
26.4742 | -35.7374 | -36.9075
26.4803 | -35.7379 | -36.9132
26.4864 | -35.7384 | -36.9185
26.4925 | -35.7389 | -36.9232
26.4986 | -35.7392 | -36.9283
26.5047 | -35.7397 | -36.9339
26.5108 | -35.7401 | -36.9388
26.5169 | -35.7405 | -36.9448
26.523 | -35.741 | -36.9492
26.5291 | -35.7414 | -36.9542
26.5352 | -35.7418 | -36.9592
26.5413 | -35.7423 | -36.9654
26.5474 | -35.7427 | -36.9708
26.5535 | -35.7431 | -36.9759
26.5596 | -35.7436 | -36.9818
26.5657 | -35.7441 | -36.9867
26.5718 | -35.7447 | -36.9924
26.578 | -35.745 | -36.9983
26.5841 | -35.7454 | -37.0039
26.5902 | -35.7459 | -37.0088
26.5963 | -35.7463 | -37.0144
26.6024 | -35.7467 | -37.0213
26.6085 | -35.7472 | -37.0259
26.6146 | -35.7476 | -37.0305
26.6207 | -35.7481 | -37.0361
26.6268 | -35.7486 | -37.043
26.6329 | -35.7491 | -37.0487
26.639 | -35.7496 | -37.054
26.6451 | -35.75 | -37.0601
26.6512 | -35.7505 | -37.0656
26.6573 | -35.7508 | -37.0717
26.6634 | -35.7512 | -37.077
26.6695 | -35.7515 | -37.0829
26.6756 | -35.752 | -37.0889
26.6817 | -35.7526 | -37.0939
26.6878 | -35.7531 | -37.1
26.6939 | -35.7536 | -37.105
26.7 | -35.7537 | -37.1102
26.7061 | -35.7542 | -37.1149
26.7122 | -35.7545 | -37.1202
26.7183 | -35.7549 | -37.1253
26.7245 | -35.7553 | -37.1311
26.7306 | -35.7558 | -37.136
26.7367 | -35.7562 | -37.1414
26.7428 | -35.7567 | -37.147
26.7489 | -35.7572 | -37.1526
26.755 | -35.7576 | -37.1583
26.7611 | -35.758 | -37.1638
26.7672 | -35.7583 | -37.1695
26.7733 | -35.7587 | -37.1743
26.7794 | -35.7591 | -37.1801
26.7855 | -35.7596 | -37.1852
26.7916 | -35.7601 | -37.1903
26.7977 | -35.7605 | -37.1961
26.8038 | -35.761 | -37.2013
26.8099 | -35.7614 | -37.2071
26.816 | -35.7617 | -37.2125
26.8221 | -35.7621 | -37.2179
26.8282 | -35.7626 | -37.2228
26.8343 | -35.763 | -37.2276
26.8404 | -35.7633 | -37.233
26.8465 | -35.7637 | -37.2386
26.8526 | -35.7641 | -37.2441
26.8587 | -35.7645 | -37.2498
26.8649 | -35.765 | -37.2555
26.871 | -35.7655 | -37.2611
26.8771 | -35.7658 | -37.2665
26.8832 | -35.7662 | -37.2716
26.8893 | -35.7667 | -37.2777
26.8954 | -35.767 | -37.2836
26.9015 | -35.7675 | -37.2891
26.9076 | -35.768 | -37.2951
26.9137 | -35.7683 | -37.3011
26.9198 | -35.7688 | -37.3069
26.9259 | -35.7692 | -37.3124
26.932 | -35.7696 | -37.3179
26.9381 | -35.77 | -37.3242
26.9442 | -35.7704 | -37.3297
26.9503 | -35.7709 | -37.3351
26.9564 | -35.7712 | -37.3406
26.9625 | -35.7718 | -37.3457
26.9686 | -35.7722 | -37.3512
26.9747 | -35.7726 | -37.3566
26.9808 | -35.7729 | -37.3622
26.9869 | -35.7733 | -37.3674
26.993 | -35.7737 | -37.372
26.9991 | -35.7741 | -37.3775
27.0052 | -35.7745 | -37.3831
27.0114 | -35.7749 | -37.3891
27.0175 | -35.7753 | -37.3948
27.0236 | -35.7758 | -37.4
27.0297 | -35.7762 | -37.4055
27.0358 | -35.7765 | -37.4109
27.0419 | -35.777 | -37.4162
27.048 | -35.7773 | -37.4216
27.0541 | -35.7777 | -37.4278
27.0602 | -35.7781 | -37.433
27.0663 | -35.7786 | -37.4388
27.0724 | -35.779 | -37.4439
27.0785 | -35.7794 | -37.4494
27.0846 | -35.7798 | -37.4548
27.0907 | -35.7802 | -37.4612
27.0968 | -35.7806 | -37.4676
27.1029 | -35.781 | -37.4726
27.109 | -35.7814 | -37.4773
27.1151 | -35.7819 | -37.4825
27.1212 | -35.7822 | -37.4867
27.1273 | -35.7826 | -37.4917
27.1334 | -35.783 | -37.4965
27.1395 | -35.7834 | -37.5015
27.1456 | -35.784 | -37.5073
27.1518 | -35.7845 | -37.5129
27.1579 | -35.785 | -37.5184
27.164 | -35.7855 | -37.5233
27.1701 | -35.7859 | -37.529
27.1762 | -35.7864 | -37.5351
27.1823 | -35.7868 | -37.5402
27.1884 | -35.7873 | -37.5462
27.1945 | -35.7876 | -37.5513
27.2006 | -35.788 | -37.557
27.2067 | -35.7885 | -37.5622
27.2128 | -35.7889 | -37.5676
27.2189 | -35.7893 | -37.5729
27.225 | -35.7897 | -37.5777
27.2311 | -35.7902 | -37.583
27.2372 | -35.7907 | -37.5883
27.2433 | -35.7911 | -37.5934
27.2494 | -35.7915 | -37.5988
27.2555 | -35.7919 | -37.6043
27.2616 | -35.7923 | -37.6103
27.2677 | -35.7927 | -37.6155
27.2738 | -35.7931 | -37.6209
27.2799 | -35.7936 | -37.6275
27.286 | -35.794 | -37.6325
27.2921 | -35.7944 | -37.6374
27.2983 | -35.7948 | -37.6426
27.3044 | -35.7953 | -37.6483
27.3105 | -35.7956 | -37.6538
27.3166 | -35.7961 | -37.6599
27.3227 | -35.7966 | -37.6652
27.3288 | -35.7969 | -37.6706
27.3349 | -35.7974 | -37.675
27.341 | -35.7978 | -37.6807
27.3471 | -35.7981 | -37.6855
27.3532 | -35.7985 | -37.6907
27.3593 | -35.7989 | -37.6963
27.3654 | -35.7993 | -37.7013
27.3715 | -35.7998 | -37.7063
27.3776 | -35.8003 | -37.7121
27.3837 | -35.8007 | -37.7179
27.3898 | -35.801 | -37.7236
27.3959 | -35.8014 | -37.7292
27.402 | -35.8019 | -37.7348
27.4081 | -35.8024 | -37.7403
27.4142 | -35.8028 | -37.7456
27.4203 | -35.8032 | -37.7506
27.4264 | -35.8036 | -37.7567
27.4325 | -35.8041 | -37.7623
27.4387 | -35.8045 | -37.768
27.4448 | -35.8049 | -37.7734
27.4509 | -35.8052 | -37.7791
27.457 | -35.8057 | -37.7843
27.4631 | -35.8061 | -37.7905
27.4692 | -35.8065 | -37.7959
27.4753 | -35.8068 | -37.7999
27.4814 | -35.8072 | -37.8053
27.4875 | -35.8075 | -37.8112
27.4936 | -35.8081 | -37.8159
27.4997 | -35.8084 | -37.8207
27.5058 | -35.8088 | -37.8256
27.5119 | -35.8093 | -37.8308
27.518 | -35.8097 | -37.8358
27.5241 | -35.81 | -37.8409
27.5302 | -35.8105 | -37.8466
27.5363 | -35.8108 | -37.8517
27.5424 | -35.8112 | -37.857
27.5485 | -35.8116 | -37.8622
27.5546 | -35.812 | -37.867
27.5607 | -35.8124 | -37.8722
27.5668 | -35.8128 | -37.8776
27.5729 | -35.8133 | -37.8825
27.5791 | -35.8136 | -37.8875
27.5852 | -35.8141 | -37.8928
27.5913 | -35.8146 | -37.8984
27.5974 | -35.8149 | -37.904
27.6035 | -35.8153 | -37.9099
27.6096 | -35.8157 | -37.915
27.6157 | -35.8161 | -37.9204
27.6218 | -35.8165 | -37.9254
27.6279 | -35.8169 | -37.931
27.634 | -35.8172 | -37.9364
27.6401 | -35.8177 | -37.9425
27.6462 | -35.8181 | -37.9483
27.6523 | -35.8185 | -37.953
27.6584 | -35.819 | -37.9582
27.6645 | -35.8194 | -37.9636
27.6706 | -35.8198 | -37.9689
27.6767 | -35.8202 | -37.9744
27.6828 | -35.8206 | -37.9792
27.6889 | -35.821 | -37.9848
27.695 | -35.8214 | -37.9908
27.7011 | -35.8218 | -37.9957
27.7072 | -35.8222 | -38.0012
27.7133 | -35.8226 | -38.0062
27.7194 | -35.823 | -38.0112
27.7256 | -35.8233 | -38.0164
27.7317 | -35.8237 | -38.0216
27.7378 | -35.8242 | -38.0271
27.7439 | -35.8246 | -38.0324
27.75 | -35.825 | -38.0385
27.7561 | -35.8254 | -38.0432
27.7622 | -35.8257 | -38.0485
27.7683 | -35.8261 | -38.054
27.7744 | -35.8264 | -38.0593
27.7805 | -35.8268 | -38.065
27.7866 | -35.8271 | -38.0706
27.7927 | -35.8275 | -38.076
27.7988 | -35.8279 | -38.0814
27.8049 | -35.8282 | -38.0867
27.811 | -35.8286 | -38.0921
27.8171 | -35.829 | -38.0974
27.8232 | -35.8294 | -38.1026
27.8293 | -35.8298 | -38.1079
27.8354 | -35.8302 | -38.1132
27.8415 | -35.8307 | -38.1183
27.8476 | -35.8311 | -38.1238
27.8537 | -35.8314 | -38.1299
27.8598 | -35.8318 | -38.1348
27.866 | -35.8322 | -38.1401
27.8721 | -35.8326 | -38.1456
27.8782 | -35.8329 | -38.1515
27.8843 | -35.8333 | -38.1568
27.8904 | -35.8337 | -38.162
27.8965 | -35.8339 | -38.1675
27.9026 | -35.8343 | -38.1728
27.9087 | -35.8347 | -38.1784
27.9148 | -35.835 | -38.1836
27.9209 | -35.8353 | -38.1894
27.927 | -35.8357 | -38.1945
27.9331 | -35.836 | -38.1999
27.9392 | -35.8363 | -38.2048
27.9453 | -35.8367 | -38.21
27.9514 | -35.8371 | -38.2157
27.9575 | -35.8375 | -38.2207
27.9636 | -35.8379 | -38.2259
27.9697 | -35.8383 | -38.2316
27.9758 | -35.8386 | -38.237
27.9819 | -35.839 | -38.2425
27.988 | -35.8394 | -38.2478
27.9941 | -35.8397 | -38.2536
28.0002 | -35.8402 | -38.2586
28.0063 | -35.8406 | -38.2641
28.0125 | -35.8409 | -38.2696
28.0186 | -35.8412 | -38.2747
28.0247 | -35.8416 | -38.2795
28.0308 | -35.842 | -38.2841
28.0369 | -35.8423 | -38.2893
28.043 | -35.8426 | -38.2951
28.0491 | -35.8429 | -38.3004
28.0552 | -35.8433 | -38.306
28.0613 | -35.8437 | -38.3112
28.0674 | -35.844 | -38.3156
28.0735 | -35.8443 | -38.3205

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4857 | -42.5495
21.4908 | -35.5494 | -35.3933
16.4948 | -34.9229 | -27.9472
12.6602 | -34.5545 | -21.6108
9.7171 | -34.3308 | -16.6712
7.4582 | -34.1736 | -12.6543
5.7244 | -34.0697 | -9.3922
4.3936 | -34.0115 | -6.7718
3.3722 | -33.9939 | -4.7034
2.5883 | -33.9911 | -3.1348
1.9866 | -33.9894 | -1.8816
1.5248 | -33.9985 | -0.87828
1.1703 | -34.0152 | -0.22387
1 | -34.0021 | 0.1016
0.89822 | -33.9789 | 0.37828
0.68941 | -33.9658 | 0.96207
0.52919 | -33.9472 | 1.4169
0.40612 | -33.966 | 2.0254
0.31171 | -33.9637 | 2.4769
0.23928 | -33.9799 | 3.2124
0.18358 | -33.9772 | 3.6459
0.14091 | -33.9924 | 4.5374
0.10819 | -34.0261 | 5.7012
0.083017 | -34.0573 | 7.186
0.063705 | -34.1192 | 9.0652
0.048923 | -34.1846 | 11.565
0.037479 | -34.2933 | 14.7645
0.028801 | -34.4601 | 18.7983
0.022078 | -34.772 | 24.1562
0.017023 | -35.2555 | 30.5716
0.013018 | -36.0682 | 38.2555
0.010014 | -37.1692 | 46.1323

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003751 | -12.3925
0.007502 | -29.004
0.011252 | -36.5641
0.015003 | -41.2199
0.018754 | -44.656
0.022505 | -46.2195
0.026255 | -47.7798
0.030006 | -49.5079
0.033757 | -50.9451
0.037508 | -53.4694
0.041258 | -53.3971
0.045009 | -54.3005
0.04876 | -54.3531
0.052511 | -56.1949
0.056261 | -57.0336
0.060012 | -57.4345
0.063763 | -59.0422
0.067514 | -60.0986
0.071264 | -60.3024
0.075015 | -60.8753
0.078766 | -60.5179
0.082517 | -60.5189
0.086268 | -61.2898
0.090018 | -62.6134
0.093769 | -63.1244
0.09752 | -63.5062
0.10127 | -65.0624
0.10502 | -65.5039
0.10877 | -66.1831
0.11252 | -65.5906
0.11627 | -65.8777
0.12002 | -67.3652
0.12377 | -67.8544
0.12753 | -67.8697
0.13128 | -67.6451
0.13503 | -66.8837
0.13878 | -68.22
0.14253 | -67.6494
0.14628 | -68.4532
0.15003 | -68.6819
0.15378 | -68.2589
0.15753 | -67.5798
0.16128 | -67.8227
0.16503 | -69.6425
0.16878 | -70.9123
0.17253 | -70.7562
0.17629 | -71.1878
0.18004 | -71.1371
0.18379 | -70.8916
0.18754 | -70.9351
0.19129 | -70.7362
0.19504 | -70.8606
0.19879 | -70.804
0.20254 | -70.7821
0.20629 | -71.0198
0.21004 | -71.5322
0.21379 | -72.4153
0.21754 | -72.7267
0.22129 | -73.178
0.22505 | -74.0017
0.2288 | -73.8651
0.23255 | -72.6708
0.2363 | -72.1103
0.24005 | -73.1935
0.2438 | -74.0744
0.24755 | -73.9127
0.2513 | -73.9696
0.25505 | -74.1272
0.2588 | -73.6629
0.26255 | -74.086
0.2663 | -75.1958
0.27005 | -75.4455
0.27381 | -76.0299
0.27756 | -74.8908
0.28131 | -74.7598
0.28506 | -75.1272
0.28881 | -75.5096
0.29256 | -76.7947
0.29631 | -77.9672
0.30006 | -77.685
0.30381 | -77.5609
0.30756 | -77.0682
0.31131 | -75.8139
0.31506 | -75.6795
0.31882 | -75.1993
0.32257 | -76.3109
0.32632 | -78.7797
0.33007 | -78.697
0.33382 | -77.7991
0.33757 | -77.267
0.34132 | -77.4735
0.34507 | -77.4453
0.34882 | -77.0187
0.35257 | -76.7166
0.35632 | -76.0689
0.36007 | -76.9681
0.36382 | -77.9856
0.36757 | -78.1206
0.37133 | -77.8615
0.37508 | -78.3601
0.37883 | -79.7733
0.38258 | -78.9487
0.38633 | -78.2336
0.39008 | -79.4824
0.39383 | -80.3247
0.39758 | -79.8279
0.40133 | -79.8972
0.40508 | -80.0601
0.40883 | -79.0383
0.41258 | -78.4271
0.41634 | -78.6876
0.42008 | -79.7623
0.42384 | -81.0731
0.42759 | -80.5088
0.43134 | -80.5038
0.43509 | -80.5086
0.43884 | -80.9766
0.44259 | -79.7822
0.44634 | -79.8876
0.45009 | -80.0444
0.45384 | -80.1263
0.45759 | -79.3849
0.46134 | -79.8078
0.46509 | -80.9125
0.46885 | -79.9677
0.4726 | -79.5759
0.47635 | -80.6951
0.4801 | -81.8391
0.48385 | -81.2447
0.4876 | -80.9028
0.49135 | -80.3512
0.4951 | -80.6972
0.49885 | -81.5954
0.5026 | -81.4418
0.50635 | -81.4706
0.5101 | -80.2493
0.51385 | -81.0169
0.5176 | -80.9963
0.52136 | -81.5207
0.52511 | -81.4226
0.52886 | -82.2128
0.53261 | -81.2427
0.53636 | -81.2679
0.54011 | -81.5236
0.54386 | -81.2685
0.54761 | -81.2471
0.55136 | -81.2582
0.55511 | -81.0271
0.55886 | -81.0392
0.56261 | -81.5385
0.56637 | -81.9579
0.57012 | -82.1581
0.57387 | -82.5056
0.57762 | -82.443
0.58137 | -81.8497
0.58512 | -81.8615
0.58887 | -82.5096
0.59262 | -82.8491
0.59637 | -82.7973
0.60012 | -82.8063
0.60387 | -82.8056
0.60762 | -82.6561
0.61137 | -82.2748
0.61513 | -82.1759
0.61888 | -82.3479
0.62263 | -82.5711
0.62638 | -82.7693
0.63013 | -82.7689
0.63388 | -83.084
0.63763 | -83.6875
0.64138 | -83.9593
0.64513 | -83.7726
0.64888 | -83.9965
0.65263 | -83.9763
0.65638 | -83.1309
0.66013 | -82.6685
0.66388 | -83.0513
0.66764 | -83.6201
0.67139 | -83.8708
0.67514 | -84.2906
0.67889 | -84.454
0.68264 | -84.2864
0.68639 | -84.3937
0.69014 | -84.6526
0.69389 | -84.9452
0.69764 | -84.8744
0.70139 | -84.3768
0.70514 | -84.117
0.70889 | -84.3772
0.71264 | -84.5379
0.7164 | -84.2631
0.72015 | -84.0672
0.7239 | -83.859
0.72765 | -83.8354
0.7314 | -84.4335
0.73515 | -85.0298
0.7389 | -85.0328
0.74265 | -85.1669
0.7464 | -85.3042
0.75015 | -85.1955
0.7539 | -84.8221
0.75765 | -84.5459
0.7614 | -84.9057
0.76516 | -85.5226
0.76891 | -85.7205
0.77266 | -85.437
0.77641 | -85.1129
0.78016 | -84.9124
0.78391 | -84.8516
0.78766 | -85.0208
0.79141 | -85.0743
0.79516 | -85.4113
0.79891 | -86.1535
0.80266 | -86.2435
0.80641 | -85.8173
0.81017 | -85.3942
0.81392 | -85.1223
0.81767 | -85.1539
0.82142 | -85.4709
0.82517 | -85.717
0.82892 | -85.9486
0.83267 | -85.8177
0.83642 | -85.5688
0.84017 | -85.6519
0.84392 | -85.7986
0.84767 | -85.9713
0.85142 | -86.2432
0.85517 | -86.1089
0.85892 | -85.6067
0.86267 | -85.3912
0.86643 | -85.801
0.87018 | -86.1545
0.87393 | -86.1498
0.87768 | -86.5194
0.88143 | -87.0475
0.88518 | -87.0691
0.88893 | -86.4807
0.89268 | -85.9297
0.89643 | -85.5617
0.90018 | -85.7522
0.90393 | -86.3407
0.90768 | -86.771
0.91143 | -87.0093
0.91519 | -86.9394
0.91894 | -87.1562
0.92269 | -87.7621
0.92644 | -87.8077
0.93019 | -87.3818
0.93394 | -86.9774
0.93769 | -86.6109
0.94144 | -86.2004
0.94519 | -85.9082
0.94894 | -86.1258
0.95269 | -86.9616
0.95644 | -87.9181
0.9602 | -87.9734
0.96395 | -87.5916
0.9677 | -87.512
0.97145 | -87.5943
0.9752 | -87.8159
0.97895 | -87.701
0.9827 | -87.4207
0.98645 | -87.489
0.9902 | -87.6962
0.99395 | -87.6317
0.9977 | -87.1314
1.0015 | -86.6402
1.0052 | -86.5346
1.009 | -86.6214
1.0127 | -87.0042
1.0165 | -87.7588
1.0202 | -88.1127
1.024 | -87.7134
1.0277 | -87.2289
1.0315 | -87.3367
1.0352 | -87.5971
1.039 | -87.5678
1.0427 | -87.7018
1.0465 | -88.4611
1.0502 | -89.0256
1.054 | -88.9513
1.0577 | -88.4179
1.0615 | -88.0073
1.0652 | -88.0177
1.069 | -87.9957
1.0727 | -87.7425
1.0765 | -87.7931
1.0802 | -87.7517
1.084 | -87.758
1.0877 | -87.8928
1.0915 | -88.0764
1.0952 | -88.1246
1.099 | -87.9077
1.1027 | -87.5456
1.1065 | -87.418
1.1102 | -87.7145
1.114 | -88.2119
1.1177 | -88.5962
1.1215 | -88.8259
1.1252 | -89.0101
1.129 | -89.112
1.1327 | -89.0475
1.1365 | -88.9077
1.1402 | -88.8871
1.144 | -88.9822
1.1477 | -89.0101
1.1515 | -88.9489
1.1552 | -88.9275
1.159 | -88.9813
1.1627 | -89.0156
1.1665 | -89.0189
1.1702 | -89.1106
1.174 | -89.2616
1.1777 | -89.3173
1.1815 | -89.2809
1.1852 | -89.2773
1.189 | -89.2851
1.1927 | -89.1442
1.1965 | -88.9475
1.2002 | -88.9344
1.204 | -89.0877
1.2077 | -89.2195
1.2115 | -89.3817
1.2152 | -89.7146
1.219 | -90.0877
1.2227 | -90.2897
1.2265 | -90.1696
1.2303 | -89.6774
1.234 | -89.2149
1.2378 | -89.1436
1.2415 | -89.3109
1.2453 | -89.402
1.249 | -89.4035
1.2528 | -89.5107
1.2565 | -89.8366
1.2603 | -90.2356
1.264 | -90.4338
1.2678 | -90.3215
1.2715 | -89.9765
1.2753 | -89.6877
1.279 | -89.6999
1.2828 | -89.961
1.2865 | -90.2317
1.2903 | -90.2722
1.294 | -90.0168
1.2978 | -89.6739
1.3015 | -89.5046
1.3053 | -89.5862
1.309 | -89.7537
1.3128 | -89.8931
1.3165 | -90.0692
1.3203 | -90.2766
1.324 | -90.409
1.3278 | -90.3736
1.3315 | -90.2593
1.3353 | -90.2577
1.339 | -90.2707
1.3428 | -90.1016
1.3465 | -89.9364
1.3503 | -89.9342
1.354 | -89.9684
1.3578 | -89.958
1.3615 | -89.9367
1.3653 | -89.9565
1.369 | -90.0449
1.3728 | -90.0429
1.3765 | -89.8095
1.3803 | -89.5876
1.384 | -89.6267
1.3878 | -89.8447
1.3915 | -90.1247
1.3953 | -90.4204
1.399 | -90.5605
1.4028 | -90.4953
1.4065 | -90.3713
1.4103 | -90.3239
1.414 | -90.4129
1.4178 | -90.6066
1.4215 | -90.7657
1.4253 | -90.7799
1.429 | -90.6864
1.4328 | -90.6388
1.4365 | -90.6835
1.4403 | -90.6248
1.444 | -90.28
1.4478 | -89.83
1.4515 | -89.6156
1.4553 | -89.7637
1.459 | -90.136
1.4628 | -90.5102
1.4665 | -90.7346
1.4703 | -90.8548
1.474 | -90.9593
1.4778 | -91.0035
1.4816 | -90.9848
1.4853 | -90.9261
1.4891 | -90.8449
1.4928 | -90.8924
1.4966 | -91.0652
1.5003 | -91.1356
1.5041 | -91.0835
1.5078 | -91.0958
1.5116 | -91.1581
1.5153 | -91.0767
1.5191 | -90.8791
1.5228 | -90.7651
1.5266 | -90.7715
1.5303 | -90.8838
1.5341 | -91.1384
1.5378 | -91.474
1.5416 | -91.7701
1.5453 | -91.8947
1.5491 | -91.8172
1.5528 | -91.6826
1.5566 | -91.5997
1.5603 | -91.5633
1.5641 | -91.6205
1.5678 | -91.7837
1.5716 | -91.8709
1.5753 | -91.7282
1.5791 | -91.5828
1.5828 | -91.6529
1.5866 | -91.8295
1.5903 | -91.9288
1.5941 | -92.0001
1.5978 | -92.1623
1.6016 | -92.3778
1.6053 | -92.4338
1.6091 | -92.2786
1.6128 | -92.1328
1.6166 | -92.1386
1.6203 | -92.1653
1.6241 | -92.1329
1.6278 | -92.0808
1.6316 | -92.0523
1.6353 | -92.1013
1.6391 | -92.1976
1.6428 | -92.2463
1.6466 | -92.178
1.6503 | -92.0863
1.6541 | -92.0901
1.6578 | -92.1711
1.6616 | -92.193
1.6653 | -92.1161
1.6691 | -92.0386
1.6728 | -92.0185
1.6766 | -92.0097
1.6803 | -91.9956
1.6841 | -91.9877
1.6878 | -91.9767
1.6916 | -91.9516
1.6953 | -91.9719
1.6991 | -92.1047
1.7028 | -92.2765
1.7066 | -92.3612
1.7103 | -92.3346
1.7141 | -92.2518
1.7178 | -92.1653
1.7216 | -92.1124
1.7253 | -92.1332
1.7291 | -92.2339
1.7329 | -92.3469
1.7366 | -92.3672
1.7404 | -92.2537
1.7441 | -92.0889
1.7479 | -92.0276
1.7516 | -92.1917
1.7554 | -92.6041
1.7591 | -93.1368
1.7629 | -93.5136
1.7666 | -93.5023
1.7704 | -93.2412
1.7741 | -93.0282
1.7779 | -92.9815
1.7816 | -93.0447
1.7854 | -93.1673
1.7891 | -93.3196
1.7929 | -93.4019
1.7966 | -93.2887
1.8004 | -93.0565
1.8041 | -92.8798
1.8079 | -92.841
1.8116 | -92.8771
1.8154 | -92.8903
1.8191 | -92.8187
1.8229 | -92.6571
1.8266 | -92.4441
1.8304 | -92.2791
1.8341 | -92.2398
1.8379 | -92.3439
1.8416 | -92.5031
1.8454 | -92.6396
1.8491 | -92.7772
1.8529 | -93.0171
1.8566 | -93.4078
1.8604 | -93.8779
1.8641 | -94.2283
1.8679 | -94.2293
1.8716 | -93.8343
1.8754 | -93.2819
1.8791 | -92.806
1.8829 | -92.4982
1.8866 | -92.3799
1.8904 | -92.4634
1.8941 | -92.728
1.8979 | -93.0413
1.9016 | -93.18
1.9054 | -93.0608
1.9091 | -92.8666
1.9129 | -92.8496
1.9166 | -93.1307
1.9204 | -93.6302
1.9241 | -94.0916
1.9279 | -94.2112
1.9316 | -93.9099
1.9354 | -93.4189
1.9391 | -92.9903
1.9429 | -92.7328
1.9466 | -92.6283
1.9504 | -92.6037
1.9541 | -92.605
1.9579 | -92.6417
1.9616 | -92.7499
1.9654 | -92.9504
1.9691 | -93.2058
1.9729 | -93.4527
1.9767 | -93.6135
1.9804 | -93.6665
1.9842 | -93.6717
1.9879 | -93.7222
1.9917 | -93.852
1.9954 | -93.996
1.9992 | -94.0342
2.0029 | -93.8794
2.0067 | -93.5864
2.0104 | -93.3575
2.0142 | -93.319
2.0179 | -93.4518
2.0217 | -93.6644
2.0254 | -93.8614
2.0292 | -93.9403
2.0329 | -93.8838
2.0367 | -93.862
2.0404 | -94.0026
2.0442 | -94.2041
2.0479 | -94.2826
2.0517 | -94.2298
2.0554 | -94.1686
2.0592 | -94.1489
2.0629 | -94.1455
2.0667 | -94.1147
2.0704 | -94.0627
2.0742 | -94.0551
2.0779 | -94.1597
2.0817 | -94.3622
2.0854 | -94.5498
2.0892 | -94.6248
2.0929 | -94.5668
2.0967 | -94.4295
2.1004 | -94.2635
2.1042 | -94.1124
2.1079 | -94.0501
2.1117 | -94.1134
2.1154 | -94.2617
2.1192 | -94.3891
2.1229 | -94.4266
2.1267 | -94.3973
2.1304 | -94.3755
2.1342 | -94.3593
2.1379 | -94.2473
2.1417 | -94.0359
2.1454 | -93.8626
2.1492 | -93.8195
2.1529 | -93.8684
2.1567 | -93.9232
2.1604 | -93.9324
2.1642 | -93.8216
2.1679 | -93.7507
2.1717 | -93.7037
2.1754 | -93.7341
2.1792 | -93.8611
2.1829 | -94.0481
2.1867 | -94.2347
2.1904 | -94.362
2.1942 | -94.4171
2.1979 | -94.445
2.2017 | -94.4777
2.2054 | -94.5289
2.2092 | -94.5936
2.213 | -94.6652
2.2167 | -94.7059
2.2205 | -94.6723
2.2242 | -94.5411
2.228 | -94.359
2.2317 | -94.1924
2.2355 | -94.058
2.2392 | -93.962
2.243 | -93.9194
2.2467 | -93.9381
2.2505 | -94.0173
2.2542 | -94.1369
2.258 | -94.2603
2.2617 | -94.3671
2.2655 | -94.4396
2.2692 | -94.4694
2.273 | -94.4843
2.2767 | -94.5072
2.2805 | -94.5357
2.2842 | -94.5577
2.288 | -94.5491
2.2917 | -94.4943
2.2955 | -94.4208
2.2992 | -94.3973
2.303 | -94.4968
2.3067 | -94.7519
2.3105 | -95.0819
2.3142 | -95.3445
2.318 | -95.3997
2.3217 | -95.2281
2.3255 | -94.9407
2.3292 | -94.6798
2.333 | -94.5428
2.3367 | -94.5552
2.3405 | -94.6876
2.3442 | -94.8403
2.348 | -94.9277
2.3517 | -94.9393
2.3555 | -94.9158
2.3592 | -94.9027
2.363 | -94.9008
2.3667 | -94.8835
2.3705 | -94.8234
2.3742 | -94.7331
2.378 | -94.6562
2.3817 | -94.6227
2.3855 | -94.6502
2.3892 | -94.735
2.393 | -94.8174
2.3967 | -94.8555
2.4005 | -94.8274
2.4042 | -94.7444
2.408 | -94.6569
2.4117 | -94.6386
2.4155 | -94.7237
2.4192 | -94.9077
2.423 | -95.1402
2.4267 | -95.332
2.4305 | -95.3949
2.4342 | -95.3324
2.438 | -95.1936
2.4417 | -95.0488
2.4455 | -94.9332
2.4492 | -94.8516
2.453 | -94.7773
2.4567 | -94.7144
2.4605 | -94.6838
2.4643 | -94.7151
2.468 | -94.8146
2.4718 | -94.9541
2.4755 | -95.0929
2.4793 | -95.1916
2.483 | -95.2264
2.4868 | -95.2197
2.4905 | -95.2199
2.4943 | -95.2661
2.498 | -95.3226
2.5018 | -95.3521
2.5055 | -95.3607
2.5093 | -95.3745
2.513 | -95.4048
2.5168 | -95.4123
2.5205 | -95.3573
2.5243 | -95.2546
2.528 | -95.1393
2.5318 | -95.0251
2.5355 | -94.9067
2.5393 | -94.8016
2.543 | -94.7409
2.5468 | -94.7621
2.5505 | -94.8849
2.5543 | -95.0752
2.558 | -95.258
2.5618 | -95.3757
2.5655 | -95.4063
2.5693 | -95.3629
2.573 | -95.3181
2.5768 | -95.3497
2.5805 | -95.4946
2.5843 | -95.7293
2.588 | -95.973
2.5918 | -96.0988
2.5955 | -96.0596
2.5993 | -95.9202
2.603 | -95.7945
2.6068 | -95.7674
2.6105 | -95.867
2.6143 | -96.037
2.618 | -96.1634
2.6218 | -96.1551
2.6255 | -96.0203
2.6293 | -95.8185
2.633 | -95.6281
2.6368 | -95.4745
2.6405 | -95.3379
2.6443 | -95.2118
2.648 | -95.0947
2.6518 | -95.0011
2.6555 | -94.955
2.6593 | -94.9868
2.663 | -95.0987
2.6668 | -95.2715
2.6705 | -95.4932
2.6743 | -95.7215
2.678 | -95.8948
2.6818 | -95.9745
2.6855 | -95.9528
2.6893 | -95.8959
2.693 | -95.849
2.6968 | -95.8476
2.7005 | -95.8878
2.7043 | -95.9747
2.7081 | -96.0474
2.7118 | -96.1255
2.7156 | -96.1877
2.7193 | -96.2168
2.7231 | -96.205
2.7268 | -96.153
2.7306 | -96.0664
2.7343 | -95.9732
2.7381 | -95.9036
2.7418 | -95.8803
2.7456 | -95.9122
2.7493 | -95.9766
2.7531 | -96.0312
2.7568 | -96.06
2.7606 | -96.0474
2.7643 | -96.0143
2.7681 | -95.9969
2.7718 | -96.0194
2.7756 | -96.0791
2.7793 | -96.1413
2.7831 | -96.1597
2.7868 | -96.1049
2.7906 | -95.9762
2.7943 | -95.8407
2.7981 | -95.7593
2.8018 | -95.771
2.8056 | -95.8918
2.8093 | -96.0942
2.8131 | -96.3149
2.8168 | -96.4888
2.8206 | -96.5766
2.8243 | -96.5706
2.8281 | -96.4985
2.8318 | -96.4005
2.8356 | -96.3191
2.8393 | -96.2644
2.8431 | -96.2363
2.8468 | -96.239
2.8506 | -96.2572
2.8543 | -96.2737
2.8581 | -96.2925
2.8618 | -96.3228
2.8656 | -96.3684
2.8693 | -96.4271
2.8731 | -96.5128
2.8768 | -96.6051
2.8806 | -96.6808
2.8843 | -96.7264
2.8881 | -96.7463
2.8918 | -96.7552
2.8956 | -96.7678
2.8993 | -96.7923
2.9031 | -96.8066
2.9068 | -96.7917
2.9106 | -96.7567
2.9143 | -96.7193
2.9181 | -96.698
2.9218 | -96.6887
2.9256 | -96.6957
2.9293 | -96.6976
2.9331 | -96.6957
2.9368 | -96.6998
2.9406 | -96.72
2.9443 | -96.7723
2.9481 | -96.851
2.9518 | -96.9238
2.9556 | -96.9658
2.9594 | -96.9682
2.9631 | -96.9464
2.9669 | -96.9296
2.9706 | -96.9311
2.9744 | -96.9537
2.9781 | -96.9937
2.9819 | -97.0262
2.9856 | -97.0321
2.9894 | -97.0152
2.9931 | -96.9895
2.9969 | -96.9576
3.0006 | -96.9458
3.0044 | -96.9523
3.0081 | -96.9762
3.0119 | -97.0185
3.0156 | -97.0597
3.0194 | -97.0883
3.0231 | -97.085
3.0269 | -97.0616
3.0306 | -97.0346
3.0344 | -97.0197
3.0381 | -97.0284
3.0419 | -97.0683
3.0456 | -97.1429
3.0494 | -97.2065
3.0531 | -97.2355
3.0569 | -97.2083
3.0606 | -97.1427
3.0644 | -97.0632
3.0681 | -96.9909
3.0719 | -96.9302
3.0756 | -96.8857
3.0794 | -96.8615
3.0831 | -96.8573
3.0869 | -96.8738
3.0906 | -96.911
3.0944 | -96.9398
3.0981 | -96.9488
3.1019 | -96.9388
3.1056 | -96.9212
3.1094 | -96.8979
3.1131 | -96.8809
3.1169 | -96.8534
3.1206 | -96.8208
3.1244 | -96.7978
3.1281 | -96.8019
3.1319 | -96.8437
3.1356 | -96.9057
3.1394 | -96.9651
3.1431 | -97.0301
3.1469 | -97.1137
3.1506 | -97.209
3.1544 | -97.3085
3.1581 | -97.3711
3.1619 | -97.3567
3.1656 | -97.2779
3.1694 | -97.1644
3.1731 | -97.0673
3.1769 | -97.038
3.1806 | -97.071
3.1844 | -97.1493
3.1881 | -97.2486
3.1919 | -97.3297
3.1956 | -97.375
3.1994 | -97.3999
3.2032 | -97.396
3.2069 | -97.3733
3.2107 | -97.3346
3.2144 | -97.2828
3.2182 | -97.2398
3.2219 | -97.215
3.2257 | -97.2173
3.2294 | -97.2491
3.2332 | -97.2997
3.2369 | -97.3584
3.2407 | -97.4294
3.2444 | -97.5306
3.2482 | -97.6302
3.2519 | -97.7365
3.2557 | -97.8454
3.2594 | -97.9256
3.2632 | -97.9655
3.2669 | -97.9672
3.2707 | -97.9332
3.2744 | -97.8748
3.2782 | -97.8099
3.2819 | -97.7508
3.2857 | -97.7107
3.2894 | -97.7168
3.2932 | -97.7582
3.2969 | -97.8116
3.3007 | -97.8372
3.3044 | -97.8215
3.3082 | -97.7773
3.3119 | -97.7273
3.3157 | -97.7073
3.3194 | -97.716
3.3232 | -97.7544
3.3269 | -97.8264
3.3307 | -97.9246
3.3344 | -98.0209
3.3382 | -98.0819
3.3419 | -98.0872
3.3457 | -98.0203
3.3494 | -97.9132
3.3532 | -97.7937
3.3569 | -97.6961
3.3607 | -97.6239
3.3644 | -97.5817
3.3682 | -97.5733
3.3719 | -97.6003
3.3757 | -97.6485
3.3794 | -97.6936
3.3832 | -97.7115
3.3869 | -97.6979
3.3907 | -97.6758
3.3944 | -97.6716
3.3982 | -97.6838
3.4019 | -97.7263
3.4057 | -97.7815
3.4094 | -97.8406
3.4132 | -97.8911
3.4169 | -97.9198
3.4207 | -97.9449
3.4244 | -97.972
3.4282 | -97.9944
3.4319 | -98.0133
3.4357 | -98.0338
3.4394 | -98.0616
3.4432 | -98.1031
3.447 | -98.1596
3.4507 | -98.2241
3.4545 | -98.2861
3.4582 | -98.3236
3.462 | -98.3277
3.4657 | -98.282
3.4695 | -98.1938
3.4732 | -98.0726
3.477 | -97.9188
3.4807 | -97.7603
3.4845 | -97.6182
3.4882 | -97.5089
3.492 | -97.4351
3.4957 | -97.4093
3.4995 | -97.4289
3.5032 | -97.4831
3.507 | -97.5776
3.5107 | -97.7015
3.5145 | -97.8345
3.5182 | -97.9732
3.522 | -98.0997
3.5257 | -98.2093
3.5295 | -98.2903
3.5332 | -98.3354
3.537 | -98.3535
3.5407 | -98.3529
3.5445 | -98.3524
3.5482 | -98.3633
3.552 | -98.3879
3.5557 | -98.4259
3.5595 | -98.4739
3.5632 | -98.5073
3.567 | -98.5267
3.5707 | -98.5413
3.5745 | -98.5444
3.5782 | -98.5237
3.582 | -98.4502
3.5857 | -98.3388
3.5895 | -98.193
3.5932 | -98.0358
3.597 | -97.9005
3.6007 | -97.8115
3.6045 | -97.7771
3.6082 | -97.8011
3.612 | -97.8813
3.6157 | -98.0031
3.6195 | -98.1502
3.6232 | -98.2882
3.627 | -98.3943
3.6307 | -98.4499
3.6345 | -98.474
3.6382 | -98.4716
3.642 | -98.4489
3.6457 | -98.4329
3.6495 | -98.4293
3.6532 | -98.4485
3.657 | -98.4881
3.6607 | -98.5376
3.6645 | -98.5764
3.6682 | -98.5687
3.672 | -98.5113
3.6757 | -98.4242
3.6795 | -98.3357
3.6832 | -98.2487
3.687 | -98.1791
3.6907 | -98.1341
3.6945 | -98.1202
3.6983 | -98.1264
3.702 | -98.1505
3.7058 | -98.1917
3.7095 | -98.22
3.7133 | -98.2347
3.717 | -98.2484
3.7208 | -98.2605
3.7245 | -98.2726
3.7283 | -98.2728
3.732 | -98.2551
3.7358 | -98.2154
3.7395 | -98.1805
3.7433 | -98.1698
3.747 | -98.1885
3.7508 | -98.2559
3.7545 | -98.3711
3.7583 | -98.5261
3.762 | -98.71
3.7658 | -98.9057
3.7695 | -99.0706
3.7733 | -99.1832
3.777 | -99.2421
3.7808 | -99.2597
3.7845 | -99.2242
3.7883 | -99.203
3.792 | -99.1551
3.7958 | -99.1136
3.7995 | -99.0652
3.8033 | -99.001
3.807 | -98.9274
3.8108 | -98.8462
3.8145 | -98.7654
3.8183 | -98.6876
3.822 | -98.6228
3.8258 | -98.5792
3.8295 | -98.5512
3.8333 | -98.5262
3.837 | -98.505
3.8408 | -98.4839
3.8445 | -98.4792
3.8483 | -98.4909
3.852 | -98.5218
3.8558 | -98.5747
3.8595 | -98.6397
3.8633 | -98.7136
3.867 | -98.7988
3.8708 | -98.8862
3.8745 | -98.9656
3.8783 | -99.035
3.882 | -99.0957
3.8858 | -99.1512
3.8895 | -99.2049
3.8933 | -99.2594
3.897 | -99.3142
3.9008 | -99.3625
3.9045 | -99.4029
3.9083 | -99.4141
3.912 | -99.4098
3.9158 | -99.3861
3.9195 | -99.3381
3.9233 | -99.2733
3.927 | -99.1981
3.9308 | -99.111
3.9345 | -99.0335
3.9383 | -98.9746
3.9421 | -98.9421
3.9458 | -98.9436
3.9496 | -98.9835
3.9533 | -99.0628
3.9571 | -99.1756
3.9608 | -99.3059
3.9646 | -99.4398
3.9683 | -99.5468
3.9721 | -99.6133
3.9758 | -99.618
3.9796 | -99.5758
3.9833 | -99.5079
3.9871 | -99.4277
3.9908 | -99.3618
3.9946 | -99.3117
3.9983 | -99.2808
4.0021 | -99.2841
4.0058 | -99.311
4.0096 | -99.36
4.0133 | -99.4135
4.0171 | -99.4546
4.0208 | -99.4726
4.0246 | -99.4664
4.0283 | -99.4225
4.0321 | -99.3425
4.0358 | -99.2412
4.0396 | -99.14
4.0433 | -99.0467
4.0471 | -98.9741
4.0508 | -98.941
4.0546 | -98.9427
4.0583 | -98.9811
4.0621 | -99.0379
4.0658 | -99.1019
4.0696 | -99.1599
4.0733 | -99.2079
4.0771 | -99.256
4.0808 | -99.3024
4.0846 | -99.3532
4.0883 | -99.4091
4.0921 | -99.4633
4.0958 | -99.5025
4.0996 | -99.5058
4.1033 | -99.4849
4.1071 | -99.447
4.1108 | -99.4077
4.1146 | -99.3727
4.1183 | -99.345
4.1221 | -99.32
4.1258 | -99.3018
4.1296 | -99.2882
4.1333 | -99.293
4.1371 | -99.3074
4.1408 | -99.3288
4.1446 | -99.3662
4.1483 | -99.4
4.1521 | -99.4341
4.1558 | -99.4822
4.1596 | -99.5391
4.1633 | -99.601
4.1671 | -99.6577
4.1708 | -99.7106
4.1746 | -99.7466
4.1783 | -99.765
4.1821 | -99.7622
4.1859 | -99.7393
4.1896 | -99.7052
4.1934 | -99.6632
4.1971 | -99.6149
4.2009 | -99.565
4.2046 | -99.5141
4.2084 | -99.4653
4.2121 | -99.4359
4.2159 | -99.4304
4.2196 | -99.4485
4.2234 | -99.4998
4.2271 | -99.583
4.2309 | -99.695
4.2346 | -99.8288
4.2384 | -99.9722
4.2421 | -100.1053
4.2459 | -100.2224
4.2496 | -100.2985
4.2534 | -100.3277
4.2571 | -100.3237
4.2609 | -100.2983
4.2646 | -100.2677
4.2684 | -100.2597
4.2721 | -100.2737
4.2759 | -100.297
4.2796 | -100.3278
4.2834 | -100.3505
4.2871 | -100.3507
4.2909 | -100.3096
4.2946 | -100.2297
4.2984 | -100.1317
4.3021 | -100.0285
4.3059 | -99.9393
4.3096 | -99.8662
4.3134 | -99.8224
4.3171 | -99.7996
4.3209 | -99.7843
4.3246 | -99.7909
4.3284 | -99.824
4.3321 | -99.8412
4.3359 | -99.857
4.3396 | -99.8698
4.3434 | -99.8826
4.3471 | -99.885
4.3509 | -99.8821
4.3546 | -99.8798
4.3584 | -99.8808
4.3621 | -99.891
4.3659 | -99.9182
4.3696 | -99.9673
4.3734 | -100.0344
4.3771 | -100.1104
4.3809 | -100.1812
4.3846 | -100.2426
4.3884 | -100.2926
4.3921 | -100.2991
4.3959 | -100.266
4.3996 | -100.2047
4.4034 | -100.1259
4.4071 | -100.0314
4.4109 | -99.9398
4.4146 | -99.8599
4.4184 | -99.7969
4.4221 | -99.7625
4.4259 | -99.7572
4.4296 | -99.7872
4.4334 | -99.8418
4.4372 | -99.9076
4.4409 | -99.9848
4.4447 | -100.0565
4.4484 | -100.1134
4.4522 | -100.1521
4.4559 | -100.1623
4.4597 | -100.1188
4.4634 | -100.039
4.4672 | -99.9336
4.4709 | -99.8072
4.4747 | -99.6812
4.4784 | -99.5753
4.4822 | -99.5026
4.4859 | -99.479
4.4897 | -99.5076
4.4934 | -99.5868
4.4972 | -99.7149
4.5009 | -99.8861
4.5047 | -100.0849
4.5084 | -100.3039
4.5122 | -100.5042
4.5159 | -100.682
4.5197 | -100.8216
4.5234 | -100.8999
4.5272 | -100.9199
4.5309 | -100.9063
4.5347 | -100.8512
4.5384 | -100.7641
4.5422 | -100.6674
4.5459 | -100.5602
4.5497 | -100.4492
4.5534 | -100.3499
4.5572 | -100.2602
4.5609 | -100.1831
4.5647 | -100.1093
4.5684 | -100.0477
4.5722 | -99.997
4.5759 | -99.9523
4.5797 | -99.9117
4.5834 | -99.8777
4.5872 | -99.8351
4.5909 | -99.7794
4.5947 | -99.7114
4.5984 | -99.6345
4.6022 | -99.5591
4.6059 | -99.4944
4.6097 | -99.4519
4.6134 | -99.4379
4.6172 | -99.4602
4.6209 | -99.522
4.6247 | -99.6095
4.6284 | -99.7222
4.6322 | -99.8519
4.6359 | -99.9929
4.6397 | -100.1275
4.6434 | -100.245
4.6472 | -100.3471
4.6509 | -100.4145
4.6547 | -100.4438
4.6584 | -100.4387
4.6622 | -100.411
4.6659 | -100.378
4.6697 | -100.3487
4.6734 | -100.3315
4.6772 | -100.3391
4.681 | -100.3629
4.6847 | -100.4094
4.6885 | -100.4673
4.6922 | -100.5256
4.696 | -100.578
4.6997 | -100.6202
4.7035 | -100.6505
4.7072 | -100.6632
4.711 | -100.6555
4.7147 | -100.6323
4.7185 | -100.6002
4.7222 | -100.577
4.726 | -100.567
4.7297 | -100.5715
4.7335 | -100.5993
4.7372 | -100.6493
4.741 | -100.7115
4.7447 | -100.7732
4.7485 | -100.8253
4.7522 | -100.8789
4.756 | -100.9127
4.7597 | -100.9363
4.7635 | -100.9537
4.7672 | -100.9644
4.771 | -100.9673
4.7747 | -100.9606
4.7785 | -100.9486
4.7822 | -100.926
4.786 | -100.8901
4.7897 | -100.8303
4.7935 | -100.7425
4.7972 | -100.6554
4.801 | -100.573
4.8047 | -100.4936
4.8085 | -100.4368
4.8122 | -100.4169
4.816 | -100.4214
4.8197 | -100.4295
4.8235 | -100.4339
4.8272 | -100.4366
4.831 | -100.4302
4.8347 | -100.4234
4.8385 | -100.4202
4.8422 | -100.4214
4.846 | -100.4372
4.8497 | -100.4645
4.8535 | -100.5043
4.8572 | -100.5529
4.861 | -100.6179
4.8647 | -100.694
4.8685 | -100.7469
4.8722 | -100.7996
4.876 | -100.8491
4.8797 | -100.8774
4.8835 | -100.8727
4.8872 | -100.8461
4.891 | -100.8029
4.8947 | -100.7464
4.8985 | -100.6746
4.9022 | -100.6002
4.906 | -100.5253
4.9097 | -100.4585
4.9135 | -100.4072
4.9172 | -100.3799
4.921 | -100.3719
4.9248 | -100.3863
4.9285 | -100.4209
4.9323 | -100.4688
4.936 | -100.5287
4.9398 | -100.6064
4.9435 | -100.6994
4.9473 | -100.8046
4.951 | -100.9032
4.9548 | -100.9876
4.9585 | -101.0576
4.9623 | -101.1007
4.966 | -101.1021
4.9698 | -101.0838
4.9735 | -101.0511
4.9773 | -101.0091
4.981 | -100.9631
4.9848 | -100.9312
4.9885 | -100.9063
4.9923 | -100.8918
4.996 | -100.8801
4.9998 | -100.8771
5.0035 | -100.8807
5.0073 | -100.8903
5.011 | -100.9048
5.0148 | -100.92
5.0185 | -100.9334
5.0223 | -100.9497
5.026 | -100.9618
5.0298 | -100.9819
5.0335 | -101.0132
5.0373 | -101.0537
5.041 | -101.0961
5.0448 | -101.1425
5.0485 | -101.1858
5.0523 | -101.2212
5.056 | -101.2449
5.0598 | -101.2546
5.0635 | -101.2368
5.0673 | -101.1885
5.071 | -101.1226
5.0748 | -101.0407
5.0785 | -100.9512
5.0823 | -100.8586
5.086 | -100.7599
5.0898 | -100.6699
5.0935 | -100.5892
5.0973 | -100.5204
5.101 | -100.4724
5.1048 | -100.4531
5.1085 | -100.4475
5.1123 | -100.4595
5.116 | -100.4892
5.1198 | -100.5405
5.1235 | -100.6137
5.1273 | -100.6967
5.131 | -100.7923
5.1348 | -100.8916
5.1385 | -100.9871
5.1423 | -101.0793
5.146 | -101.1629
5.1498 | -101.2236
5.1535 | -101.2588
5.1573 | -101.275
5.161 | -101.2813
5.1648 | -101.2895
5.1685 | -101.2917
5.1723 | -101.2953
5.1761 | -101.3128
5.1798 | -101.3271
5.1836 | -101.3493
5.1873 | -101.3944
5.1911 | -101.4383
5.1948 | -101.4712
5.1986 | -101.4862
5.2023 | -101.4758
5.2061 | -101.4585
5.2098 | -101.4261
5.2136 | -101.3778
5.2173 | -101.3244
5.2211 | -101.2697
5.2248 | -101.2198
5.2286 | -101.1852
5.2323 | -101.1655
5.2361 | -101.1517
5.2398 | -101.1482
5.2436 | -101.1559
5.2473 | -101.176
5.2511 | -101.2072
5.2548 | -101.239
5.2586 | -101.275
5.2623 | -101.3042
5.2661 | -101.3223
5.2698 | -101.334
5.2736 | -101.3331
5.2773 | -101.3192
5.2811 | -101.3059
5.2848 | -101.293
5.2886 | -101.2901
5.2923 | -101.3008
5.2961 | -101.3319
5.2998 | -101.3809
5.3036 | -101.441
5.3073 | -101.4992
5.3111 | -101.5595
5.3148 | -101.6061
5.3186 | -101.6295
5.3223 | -101.6353
5.3261 | -101.6166
5.3298 | -101.5802
5.3336 | -101.535
5.3373 | -101.4815
5.3411 | -101.4235
5.3448 | -101.3687
5.3486 | -101.3235
5.3523 | -101.2874
5.3561 | -101.2559
5.3598 | -101.2266
5.3636 | -101.2117
5.3673 | -101.2132
5.3711 | -101.2275
5.3748 | -101.2436
5.3786 | -101.2607
5.3823 | -101.2814
5.3861 | -101.2962
5.3898 | -101.2933
5.3936 | -101.273
5.3973 | -101.24
5.4011 | -101.1936
5.4048 | -101.1435
5.4086 | -101.1104
5.4123 | -101.0812
5.4161 | -101.0687
5.4199 | -101.0774
5.4236 | -101.1056
5.4274 | -101.1517
5.4311 | -101.2171
5.4349 | -101.2998
5.4386 | -101.391
5.4424 | -101.4839
5.4461 | -101.5612
5.4499 | -101.6203
5.4536 | -101.6595
5.4574 | -101.6801
5.4611 | -101.6822
5.4649 | -101.6636
5.4686 | -101.641
5.4724 | -101.6206
5.4761 | -101.6047
5.4799 | -101.5915
5.4836 | -101.5812
5.4874 | -101.5804
5.4911 | -101.5886
5.4949 | -101.5995
5.4986 | -101.6147
5.5024 | -101.6261
5.5061 | -101.6372
5.5099 | -101.6412
5.5136 | -101.6376
5.5174 | -101.6304
5.5211 | -101.6153
5.5249 | -101.6019
5.5286 | -101.5857
5.5324 | -101.5842
5.5361 | -101.5919
5.5399 | -101.605
5.5436 | -101.6249
5.5474 | -101.6487
5.5511 | -101.6714
5.5549 | -101.6853
5.5586 | -101.6915
5.5624 | -101.6932
5.5661 | -101.6872
5.5699 | -101.671
5.5736 | -101.6591
5.5774 | -101.646
5.5811 | -101.6324
5.5849 | -101.6188
5.5886 | -101.6091
5.5924 | -101.6041
5.5961 | -101.6054
5.5999 | -101.6123
5.6036 | -101.6178
5.6074 | -101.6372
5.6111 | -101.6673
5.6149 | -101.7064
5.6186 | -101.7504
5.6224 | -101.8092
5.6261 | -101.8756
5.6299 | -101.9324
5.6336 | -101.9839
5.6374 | -102.0201
5.6411 | -102.0394
5.6449 | -102.0365
5.6486 | -102.0141
5.6524 | -101.9675
5.6561 | -101.912
5.6599 | -101.8417
5.6636 | -101.7683
5.6674 | -101.7023
5.6712 | -101.6497
5.6749 | -101.6018
5.6787 | -101.5718
5.6824 | -101.56
5.6862 | -101.5649
5.6899 | -101.5927
5.6937 | -101.6322
5.6974 | -101.686
5.7012 | -101.7526
5.7049 | -101.8216
5.7087 | -101.8913
5.7124 | -101.9599
5.7162 | -102.0275
5.7199 | -102.0821
5.7237 | -102.1346
5.7274 | -102.179
5.7312 | -102.2178
5.7349 | -102.2533
5.7387 | -102.2845
5.7424 | -102.302
5.7462 | -102.3051
5.7499 | -102.3006
5.7537 | -102.2794
5.7574 | -102.2371
5.7612 | -102.1828
5.7649 | -102.1175
5.7687 | -102.049
5.7724 | -101.9816
5.7762 | -101.9097
5.7799 | -101.8439
5.7837 | -101.7836
5.7874 | -101.7367
5.7912 | -101.7082
5.7949 | -101.695
5.7987 | -101.6991
5.8024 | -101.7272
5.8062 | -101.7681
5.8099 | -101.816
5.8137 | -101.8676
5.8174 | -101.9149
5.8212 | -101.9584
5.8249 | -101.9978
5.8287 | -102.0249
5.8324 | -102.0401
5.8362 | -102.04
5.8399 | -102.0361
5.8437 | -102.0381
5.8474 | -102.0376
5.8512 | -102.0362
5.8549 | -102.0352
5.8587 | -102.0267
5.8624 | -102.0116
5.8662 | -101.9855
5.8699 | -101.9509
5.8737 | -101.9064
5.8774 | -101.8499
5.8812 | -101.7949
5.8849 | -101.7466
5.8887 | -101.6928
5.8924 | -101.6501
5.8962 | -101.6178
5.8999 | -101.5923
5.9037 | -101.5738
5.9074 | -101.5549
5.9112 | -101.5461
5.915 | -101.5524
5.9187 | -101.5687
5.9225 | -101.5999
5.9262 | -101.6413
5.93 | -101.6841
5.9337 | -101.7266
5.9375 | -101.7592
5.9412 | -101.786
5.945 | -101.8008
5.9487 | -101.7838
5.9525 | -101.7736
5.9562 | -101.7543
5.96 | -101.7309
5.9637 | -101.7045
5.9675 | -101.6813
5.9712 | -101.6612
5.975 | -101.6443
5.9787 | -101.6382
5.9825 | -101.6339
5.9862 | -101.6287
5.99 | -101.6327
5.9937 | -101.6479
5.9975 | -101.671
6.0012 | -101.7048
6.005 | -101.7485
6.0087 | -101.8037
6.0125 | -101.8663
6.0162 | -101.9309
6.02 | -102.0032
6.0237 | -102.0872
6.0275 | -102.1686
6.0312 | -102.2357
6.035 | -102.2972
6.0387 | -102.3469
6.0425 | -102.3914
6.0462 | -102.4255
6.05 | -102.4507
6.0537 | -102.4654
6.0575 | -102.4688
6.0612 | -102.4559
6.065 | -102.4294
6.0687 | -102.4055
6.0725 | -102.3834
6.0762 | -102.3567
6.08 | -102.3289
6.0837 | -102.2993
6.0875 | -102.2707
6.0912 | -102.2465
6.095 | -102.2209
6.0987 | -102.1979
6.1025 | -102.1848
6.1062 | -102.1756
6.11 | -102.1767
6.1137 | -102.1883
6.1175 | -102.2032
6.1212 | -102.2309
6.125 | -102.2695
6.1287 | -102.3118
6.1325 | -102.3537
6.1362 | -102.3944
6.14 | -102.4327
6.1437 | -102.4602
6.1475 | -102.4895
6.1512 | -102.5264
6.155 | -102.5713
6.1588 | -102.6144
6.1625 | -102.6537
6.1663 | -102.6909
6.17 | -102.7313
6.1738 | -102.7653
6.1775 | -102.7956
6.1813 | -102.8173
6.185 | -102.8304
6.1888 | -102.8352
6.1925 | -102.8172
6.1963 | -102.7797
6.2 | -102.7376
6.2038 | -102.6788
6.2075 | -102.6036
6.2113 | -102.5338
6.215 | -102.4697
6.2188 | -102.4138
6.2225 | -102.3695
6.2263 | -102.3406
6.23 | -102.329
6.2338 | -102.3433
6.2375 | -102.3759
6.2413 | -102.4154
6.245 | -102.4639
6.2488 | -102.5227
6.2525 | -102.5807
6.2563 | -102.6286
6.26 | -102.6685
6.2638 | -102.6981
6.2675 | -102.703
6.2713 | -102.6768
6.275 | -102.634
6.2788 | -102.5737
6.2825 | -102.5049
6.2863 | -102.4451
6.29 | -102.3886
6.2938 | -102.3353
6.2975 | -102.3006
6.3013 | -102.288
6.305 | -102.2882
6.3088 | -102.3007
6.3125 | -102.3266
6.3163 | -102.3622
6.32 | -102.3995
6.3238 | -102.4339
6.3275 | -102.465
6.3313 | -102.4929
6.335 | -102.5096
6.3388 | -102.5211
6.3425 | -102.5274
6.3463 | -102.5277
6.35 | -102.522
6.3538 | -102.5045
6.3575 | -102.4764
6.3613 | -102.447
6.365 | -102.4183
6.3688 | -102.3807
6.3725 | -102.3408
6.3763 | -102.2998
6.38 | -102.2521
6.3838 | -102.2077
6.3875 | -102.1743
6.3913 | -102.1476
6.395 | -102.1305
6.3988 | -102.1278
6.4026 | -102.1384
6.4063 | -102.162
6.4101 | -102.189
6.4138 | -102.2235
6.4176 | -102.2688
6.4213 | -102.3131
6.4251 | -102.36
6.4288 | -102.4177
6.4326 | -102.4686
6.4363 | -102.5102
6.4401 | -102.5551
6.4438 | -102.5942
6.4476 | -102.6288
6.4513 | -102.6621
6.4551 | -102.6933
6.4588 | -102.7214
6.4626 | -102.7582
6.4663 | -102.7982
6.4701 | -102.8336
6.4738 | -102.8724
6.4776 | -102.9105
6.4813 | -102.9366
6.4851 | -102.9588
6.4888 | -102.9757
6.4926 | -102.9801
6.4963 | -102.9831
6.5001 | -102.9848
6.5038 | -102.9856
6.5076 | -102.994
6.5113 | -102.998
6.5151 | -103
6.5188 | -103.0082
6.5226 | -103.0186
6.5263 | -103.0294
6.5301 | -103.0474
6.5338 | -103.0697
6.5376 | -103.0906
6.5413 | -103.1104
6.5451 | -103.1274
6.5488 | -103.1387
6.5526 | -103.1503
6.5563 | -103.1555
6.5601 | -103.1496
6.5638 | -103.1402
6.5676 | -103.1264
6.5713 | -103.1006
6.5751 | -103.0662
6.5788 | -103.0303
6.5826 | -102.9875
6.5863 | -102.9354
6.5901 | -102.8924
6.5938 | -102.8542
6.5976 | -102.8198
6.6013 | -102.7893
6.6051 | -102.7615
6.6088 | -102.7348
6.6126 | -102.712
6.6163 | -102.6899
6.6201 | -102.6744
6.6238 | -102.6649
6.6276 | -102.6581
6.6313 | -102.6573
6.6351 | -102.6653
6.6388 | -102.6769
6.6426 | -102.6987
6.6463 | -102.7317
6.6501 | -102.7718
6.6539 | -102.8125
6.6576 | -102.8631
6.6614 | -102.9228
6.6651 | -102.9834
6.6689 | -103.0449
6.6726 | -103.109
6.6764 | -103.168
6.6801 | -103.2208
6.6839 | -103.2709
6.6876 | -103.3163
6.6914 | -103.3376
6.6951 | -103.3444
6.6989 | -103.347
6.7026 | -103.3383
6.7064 | -103.3226
6.7101 | -103.3012
6.7139 | -103.2784
6.7176 | -103.2519
6.7214 | -103.2188
6.7251 | -103.1881
6.7289 | -103.1571
6.7326 | -103.1269
6.7364 | -103.0959
6.7401 | -103.068
6.7439 | -103.0408
6.7476 | -103.0105
6.7514 | -102.986
6.7551 | -102.9742
6.7589 | -102.9635
6.7626 | -102.9604
6.7664 | -102.9691
6.7701 | -102.986
6.7739 | -103.0112
6.7776 | -103.043
6.7814 | -103.0838
6.7851 | -103.1316
6.7889 | -103.1864
6.7926 | -103.2436
6.7964 | -103.3029
6.8001 | -103.3492
6.8039 | -103.3857
6.8076 | -103.416
6.8114 | -103.4397
6.8151 | -103.4551
6.8189 | -103.4697
6.8226 | -103.4836
6.8264 | -103.4951
6.8301 | -103.5045
6.8339 | -103.5147
6.8376 | -103.5259
6.8414 | -103.5345
6.8451 | -103.5381
6.8489 | -103.5379
6.8526 | -103.529
6.8564 | -103.5109
6.8601 | -103.4948
6.8639 | -103.474
6.8676 | -103.4479
6.8714 | -103.4243
6.8751 | -103.3979
6.8789 | -103.3722
6.8826 | -103.3541
6.8864 | -103.3419
6.8902 | -103.3402
6.8939 | -103.3453
6.8977 | -103.3488
6.9014 | -103.3653
6.9052 | -103.3848
6.9089 | -103.3989
6.9127 | -103.4103
6.9164 | -103.4166
6.9202 | -103.41
6.9239 | -103.3989
6.9277 | -103.3861
6.9314 | -103.3708
6.9352 | -103.3472
6.9389 | -103.3162
6.9427 | -103.285
6.9464 | -103.2517
6.9502 | -103.2221
6.9539 | -103.1904
6.9577 | -103.1633
6.9614 | -103.1403
6.9652 | -103.1268
6.9689 | -103.1207
6.9727 | -103.1098
6.9764 | -103.0999
6.9802 | -103.097
6.9839 | -103.0852
6.9877 | -103.0682
6.9914 | -103.0469
6.9952 | -103.0274
6.9989 | -103.0115
7.0027 | -102.988
7.0064 | -102.9669
7.0102 | -102.9551
7.0139 | -102.9486
7.0177 | -102.9509
7.0214 | -102.9668
7.0252 | -102.9841
7.0289 | -103.0013
7.0327 | -103.036
7.0364 | -103.0639
7.0402 | -103.0988
7.0439 | -103.1365
7.0477 | -103.1717
7.0514 | -103.2108
7.0552 | -103.249
7.0589 | -103.2905
7.0627 | -103.3285
7.0664 | -103.3642
7.0702 | -103.3942
7.0739 | -103.4256
7.0777 | -103.4664
7.0814 | -103.5036
7.0852 | -103.5366
7.0889 | -103.5701
7.0927 | -103.6004
7.0964 | -103.6242
7.1002 | -103.6421
7.1039 | -103.6543
7.1077 | -103.6601
7.1114 | -103.6589
7.1152 | -103.6489
7.1189 | -103.6327
7.1227 | -103.6116
7.1264 | -103.5806
7.1302 | -103.5538
7.1339 | -103.5241
7.1377 | -103.4879
7.1415 | -103.4446
7.1452 | -103.3964
7.149 | -103.341
7.1527 | -103.2867
7.1565 | -103.2376
7.1602 | -103.1884
7.164 | -103.1464
7.1677 | -103.1076
7.1715 | -103.0747
7.1752 | -103.0533
7.179 | -103.0379
7.1827 | -103.0288
7.1865 | -103.0244
7.1902 | -103.0282
7.194 | -103.0378
7.1977 | -103.0598
7.2015 | -103.0896
7.2052 | -103.1209
7.209 | -103.1512
7.2127 | -103.1791
7.2165 | -103.2094
7.2202 | -103.2369
7.224 | -103.2592
7.2277 | -103.2756
7.2315 | -103.2884
7.2352 | -103.2994
7.239 | -103.3065
7.2427 | -103.3124
7.2465 | -103.3197
7.2502 | -103.3352
7.254 | -103.3563
7.2577 | -103.3805
7.2615 | -103.4043
7.2652 | -103.4241
7.269 | -103.4416
7.2727 | -103.4578
7.2765 | -103.4796
7.2802 | -103.4995
7.284 | -103.514
7.2877 | -103.5339
7.2915 | -103.5517
7.2952 | -103.567
7.299 | -103.5847
7.3027 | -103.5999
7.3065 | -103.6119
7.3102 | -103.6279
7.314 | -103.6475
7.3177 | -103.668
7.3215 | -103.6842
7.3252 | -103.6896
7.329 | -103.6865
7.3327 | -103.6763
7.3365 | -103.6666
7.3402 | -103.6465
7.344 | -103.6201
7.3477 | -103.5892
7.3515 | -103.5618
7.3552 | -103.5324
7.359 | -103.4887
7.3627 | -103.4472
7.3665 | -103.417
7.3702 | -103.3958
7.374 | -103.3776
7.3777 | -103.3653
7.3815 | -103.3541
7.3853 | -103.3437
7.389 | -103.3393
7.3928 | -103.3408
7.3965 | -103.3439
7.4003 | -103.3532
7.404 | -103.3673
7.4078 | -103.3839
7.4115 | -103.4079
7.4153 | -103.4305
7.419 | -103.4504
7.4228 | -103.4768
7.4265 | -103.5004
7.4303 | -103.5184
7.434 | -103.5309
7.4378 | -103.5393
7.4415 | -103.5391
7.4453 | -103.5223
7.449 | -103.5004
7.4528 | -103.4779
7.4565 | -103.4548
7.4603 | -103.4229
7.464 | -103.397
7.4678 | -103.3858
7.4715 | -103.3757
7.4753 | -103.3646
7.479 | -103.3573
7.4828 | -103.3608
7.4865 | -103.3692
7.4903 | -103.3787
7.494 | -103.3924
7.4978 | -103.4064
7.5015 | -103.4244
7.5053 | -103.4457
7.509 | -103.4627
7.5128 | -103.4797
7.5165 | -103.497
7.5203 | -103.5118
7.524 | -103.5307
7.5278 | -103.5476
7.5315 | -103.5584
7.5353 | -103.5767
7.539 | -103.5949
7.5428 | -103.6174
7.5465 | -103.6427
7.5503 | -103.6737
7.554 | -103.7137
7.5578 | -103.7597
7.5615 | -103.8043
7.5653 | -103.8494
7.569 | -103.9014
7.5728 | -103.9374
7.5765 | -103.9719
7.5803 | -103.9981
7.584 | -104.013
7.5878 | -104.0144
7.5915 | -104.0081
7.5953 | -103.9939
7.599 | -103.9675
7.6028 | -103.9319
7.6065 | -103.8873
7.6103 | -103.8335
7.614 | -103.7697
7.6178 | -103.71
7.6215 | -103.6521
7.6253 | -103.5947
7.6291 | -103.5391
7.6328 | -103.4888
7.6366 | -103.4457
7.6403 | -103.4168
7.6441 | -103.3961
7.6478 | -103.3846
7.6516 | -103.3856
7.6553 | -103.3907
7.6591 | -103.4069
7.6628 | -103.4317
7.6666 | -103.4597
7.6703 | -103.4936
7.6741 | -103.5315
7.6778 | -103.5714
7.6816 | -103.6106
7.6853 | -103.6497
7.6891 | -103.6897
7.6928 | -103.7328
7.6966 | -103.7741
7.7003 | -103.8178
7.7041 | -103.8666
7.7078 | -103.9117
7.7116 | -103.9509
7.7153 | -103.9963
7.7191 | -104.0423
7.7228 | -104.0833
7.7266 | -104.1212
7.7303 | -104.1519
7.7341 | -104.1752
7.7378 | -104.2026
7.7416 | -104.227
7.7453 | -104.246
7.7491 | -104.2619
7.7528 | -104.2745
7.7566 | -104.2811
7.7603 | -104.2835
7.7641 | -104.2832
7.7678 | -104.2826
7.7716 | -104.2743
7.7753 | -104.2594
7.7791 | -104.2418
7.7828 | -104.2293
7.7866 | -104.2169
7.7903 | -104.2045
7.7941 | -104.1951
7.7978 | -104.187
7.8016 | -104.1785
7.8053 | -104.1808
7.8091 | -104.1826
7.8128 | -104.1864
7.8166 | -104.1894
7.8203 | -104.1859
7.8241 | -104.1841
7.8278 | -104.1846
7.8316 | -104.1818
7.8353 | -104.1765
7.8391 | -104.168
7.8428 | -104.1577
7.8466 | -104.1404
7.8503 | -104.1161
7.8541 | -104.0833
7.8578 | -104.0537
7.8616 | -104.0287
7.8653 | -104.0044
7.8691 | -103.9821
7.8728 | -103.9601
7.8766 | -103.9456
7.8804 | -103.9343
7.8841 | -103.9332
7.8879 | -103.9399
7.8916 | -103.9483
7.8954 | -103.9601
7.8991 | -103.9709
7.9029 | -103.9851
7.9066 | -104.0046
7.9104 | -104.0192
7.9141 | -104.03
7.9179 | -104.0355
7.9216 | -104.0279
7.9254 | -104.0267
7.9291 | -104.0224
7.9329 | -104.0171
7.9366 | -104.011
7.9404 | -104.0003
7.9441 | -103.9905
7.9479 | -103.9851
7.9516 | -103.9867
7.9554 | -103.995
7.9591 | -103.9996
7.9629 | -104.0098
7.9666 | -104.0245
7.9704 | -104.0396
7.9741 | -104.0505
7.9779 | -104.0631
7.9816 | -104.0767
7.9854 | -104.0923
7.9891 | -104.1106
7.9929 | -104.1244
7.9966 | -104.1301
8.0004 | -104.128
8.0041 | -104.1289
8.0079 | -104.1274
8.0116 | -104.1263
8.0154 | -104.1244
8.0191 | -104.116
8.0229 | -104.1049
8.0266 | -104.1011
8.0304 | -104.1006
8.0341 | -104.1026
8.0379 | -104.1039
8.0416 | -104.1062
8.0454 | -104.1141
8.0491 | -104.1228
8.0529 | -104.1367
8.0566 | -104.1565
8.0604 | -104.1729
8.0641 | -104.1858
8.0679 | -104.1953
8.0716 | -104.2004
8.0754 | -104.2116
8.0791 | -104.2117
8.0829 | -104.1989
8.0866 | -104.1792
8.0904 | -104.1519
8.0941 | -104.117
8.0979 | -104.0787
8.1016 | -104.0437
8.1054 | -104.0053
8.1091 | -103.9678
8.1129 | -103.9361
8.1166 | -103.9089
8.1204 | -103.8823
8.1242 | -103.8544
8.1279 | -103.831
8.1317 | -103.8091
8.1354 | -103.7924
8.1392 | -103.7866
8.1429 | -103.7854
8.1467 | -103.7852
8.1504 | -103.7908
8.1542 | -103.8036
8.1579 | -103.8232
8.1617 | -103.8431
8.1654 | -103.865
8.1692 | -103.8895
8.1729 | -103.9149
8.1767 | -103.9372
8.1804 | -103.9617
8.1842 | -103.9867
8.1879 | -104.0131
8.1917 | -104.0373
8.1954 | -104.0583
8.1992 | -104.0828
8.2029 | -104.1048
8.2067 | -104.1271
8.2104 | -104.1532
8.2142 | -104.1773
8.2179 | -104.2023
8.2217 | -104.2245
8.2254 | -104.2531
8.2292 | -104.2882
8.2329 | -104.3272
8.2367 | -104.3692
8.2404 | -104.4105
8.2442 | -104.4556
8.2479 | -104.5023
8.2517 | -104.545
8.2554 | -104.5891
8.2592 | -104.6287
8.2629 | -104.6597
8.2667 | -104.6886
8.2704 | -104.7079
8.2742 | -104.7177
8.2779 | -104.7177
8.2817 | -104.7099
8.2854 | -104.6915
8.2892 | -104.6629
8.2929 | -104.6259
8.2967 | -104.578
8.3004 | -104.5246
8.3042 | -104.4704
8.3079 | -104.4105
8.3117 | -104.3435
8.3154 | -104.2793
8.3192 | -104.2162
8.3229 | -104.1557
8.3267 | -104.0994
8.3304 | -104.0504
8.3342 | -103.9993
8.3379 | -103.9511
8.3417 | -103.9082
8.3454 | -103.8729
8.3492 | -103.8426
8.3529 | -103.8178
8.3567 | -103.7975
8.3604 | -103.7871
8.3642 | -103.7859
8.368 | -103.7871
8.3717 | -103.7934
8.3755 | -103.8057
8.3792 | -103.8175
8.383 | -103.8314
8.3867 | -103.8508
8.3905 | -103.8694
8.3942 | -103.888
8.398 | -103.9068
8.4017 | -103.9233
8.4055 | -103.9366
8.4092 | -103.9471
8.413 | -103.9606
8.4167 | -103.9712
8.4205 | -103.9807
8.4242 | -103.9902
8.428 | -103.9942
8.4317 | -103.9942
8.4355 | -103.9968
8.4392 | -103.9991
8.443 | -104.0001
8.4467 | -104.0013
8.4505 | -104.0019
8.4542 | -104.001
8.458 | -104.0033
8.4617 | -104.013
8.4655 | -104.0257
8.4692 | -104.0406
8.473 | -104.0603
8.4767 | -104.0853
8.4805 | -104.1182
8.4842 | -104.1539
8.488 | -104.198
8.4917 | -104.2508
8.4955 | -104.3052
8.4992 | -104.3624
8.503 | -104.4253
8.5067 | -104.4839
8.5105 | -104.5346
8.5142 | -104.5848
8.518 | -104.6314
8.5217 | -104.6687
8.5255 | -104.6961
8.5292 | -104.7123
8.533 | -104.7221
8.5367 | -104.713
8.5405 | -104.6943
8.5442 | -104.6708
8.548 | -104.6315
8.5517 | -104.58
8.5555 | -104.5263
8.5592 | -104.4664
8.563 | -104.4028
8.5667 | -104.3365
8.5705 | -104.2656
8.5742 | -104.1962
8.578 | -104.1301
8.5817 | -104.0698
8.5855 | -104.0082
8.5892 | -103.9562
8.593 | -103.9157
8.5967 | -103.8873
8.6005 | -103.8756
8.6042 | -103.8799
8.608 | -103.8993
8.6117 | -103.9291
8.6155 | -103.9715
8.6193 | -104.0316
8.623 | -104.1022
8.6268 | -104.18
8.6305 | -104.2636
8.6343 | -104.3478
8.638 | -104.4317
8.6418 | -104.5108
8.6455 | -104.5851
8.6493 | -104.6493
8.653 | -104.6987
8.6568 | -104.7087
8.6605 | -104.7339
8.6643 | -104.7573
8.668 | -104.7735
8.6718 | -104.7748
8.6755 | -104.7744
8.6793 | -104.7751
8.683 | -104.7692
8.6868 | -104.7511
8.6905 | -104.7257
8.6943 | -104.7004
8.698 | -104.6811
8.7018 | -104.6641
8.7055 | -104.6471
8.7093 | -104.6313
8.713 | -104.6214
8.7168 | -104.604
8.7205 | -104.589
8.7243 | -104.5763
8.728 | -104.5613
8.7318 | -104.545
8.7355 | -104.5321
8.7393 | -104.5219
8.743 | -104.505
8.7468 | -104.4864
8.7505 | -104.4666
8.7543 | -104.443
8.758 | -104.4232
8.7618 | -104.4041
8.7655 | -104.3817
8.7693 | -104.3617
8.773 | -104.3446
8.7768 | -104.3306
8.7805 | -104.3255
8.7843 | -104.3172
8.788 | -104.3167
8.7918 | -104.3271
8.7955 | -104.3383
8.7993 | -104.3555
8.803 | -104.3824
8.8068 | -104.4114
8.8105 | -104.4404
8.8143 | -104.4761
8.818 | -104.5197
8.8218 | -104.5628
8.8255 | -104.6045
8.8293 | -104.6377
8.833 | -104.6644
8.8368 | -104.6866
8.8405 | -104.705
8.8443 | -104.7178
8.848 | -104.727
8.8518 | -104.7261
8.8555 | -104.7189
8.8593 | -104.7054
8.863 | -104.6893
8.8668 | -104.6722
8.8706 | -104.6551
8.8743 | -104.6395
8.8781 | -104.6264
8.8818 | -104.6229
8.8856 | -104.6235
8.8893 | -104.6296
8.8931 | -104.6396
8.8968 | -104.6549
8.9006 | -104.675
8.9043 | -104.6946
8.9081 | -104.7196
8.9118 | -104.743
8.9156 | -104.7636
8.9193 | -104.7812
8.9231 | -104.7945
8.9268 | -104.8024
8.9306 | -104.8091
8.9343 | -104.8092
8.9381 | -104.8042
8.9418 | -104.799
8.9456 | -104.7844
8.9493 | -104.7643
8.9531 | -104.7462
8.9568 | -104.7266
8.9606 | -104.7069
8.9643 | -104.6862
8.9681 | -104.6567
8.9718 | -104.6257
8.9756 | -104.5974
8.9793 | -104.5738
8.9831 | -104.5525
8.9868 | -104.5314
8.9906 | -104.5116
8.9943 | -104.4943
8.9981 | -104.4771
9.0018 | -104.4646
9.0056 | -104.4513
9.0093 | -104.4337
9.0131 | -104.4151
9.0168 | -104.4021
9.0206 | -104.3882
9.0243 | -104.3815
9.0281 | -104.3781
9.0318 | -104.3726
9.0356 | -104.3669
9.0393 | -104.3604
9.0431 | -104.3552
9.0468 | -104.3513
9.0506 | -104.3477
9.0543 | -104.343
9.0581 | -104.3384
9.0618 | -104.3343
9.0656 | -104.3338
9.0693 | -104.3352
9.0731 | -104.3352
9.0768 | -104.3413
9.0806 | -104.3535
9.0843 | -104.3666
9.0881 | -104.3784
9.0918 | -104.3964
9.0956 | -104.4149
9.0993 | -104.4441
9.1031 | -104.4717
9.1068 | -104.4988
9.1106 | -104.5244
9.1144 | -104.5489
9.1181 | -104.5787
9.1219 | -104.611
9.1256 | -104.6436
9.1294 | -104.676
9.1331 | -104.7071
9.1369 | -104.7452
9.1406 | -104.7835
9.1444 | -104.8135
9.1481 | -104.8407
9.1519 | -104.8646
9.1556 | -104.8804
9.1594 | -104.8922
9.1631 | -104.9038
9.1669 | -104.9064
9.1706 | -104.904
9.1744 | -104.8989
9.1781 | -104.8883
9.1819 | -104.8683
9.1856 | -104.8454
9.1894 | -104.8212
9.1931 | -104.7922
9.1969 | -104.7529
9.2006 | -104.7244
9.2044 | -104.6972
9.2081 | -104.6687
9.2119 | -104.6378
9.2156 | -104.6098
9.2194 | -104.5849
9.2231 | -104.5672
9.2269 | -104.5533
9.2306 | -104.5444
9.2344 | -104.5375
9.2381 | -104.5373
9.2419 | -104.5439
9.2456 | -104.5569
9.2494 | -104.5726
9.2531 | -104.5902
9.2569 | -104.6128
9.2606 | -104.639
9.2644 | -104.6659
9.2681 | -104.6946
9.2719 | -104.7222
9.2756 | -104.7453
9.2794 | -104.7716
9.2831 | -104.7956
9.2869 | -104.8197
9.2906 | -104.8408
9.2944 | -104.861
9.2981 | -104.8748
9.3019 | -104.8867
9.3056 | -104.8907
9.3094 | -104.8904
9.3131 | -104.8925
9.3169 | -104.8911
9.3206 | -104.885
9.3244 | -104.8725
9.3281 | -104.8544
9.3319 | -104.8374
9.3356 | -104.8157
9.3394 | -104.7887
9.3431 | -104.7664
9.3469 | -104.7441
9.3506 | -104.7222
9.3544 | -104.7013
9.3582 | -104.6823
9.3619 | -104.6636
9.3657 | -104.6466
9.3694 | -104.6344
9.3732 | -104.6281
9.3769 | -104.6307
9.3807 | -104.632
9.3844 | -104.6406
9.3882 | -104.6603
9.3919 | -104.6825
9.3957 | -104.705
9.3994 | -104.734
9.4032 | -104.7645
9.4069 | -104.7937
9.4107 | -104.8256
9.4144 | -104.8556
9.4182 | -104.8805
9.4219 | -104.9043
9.4257 | -104.9194
9.4294 | -104.932
9.4332 | -104.9386
9.4369 | -104.9387
9.4407 | -104.9321
9.4444 | -104.92
9.4482 | -104.8974
9.4519 | -104.8694
9.4557 | -104.8419
9.4594 | -104.8059
9.4632 | -104.7717
9.4669 | -104.7395
9.4707 | -104.7066
9.4744 | -104.6747
9.4782 | -104.6469
9.4819 | -104.6176
9.4857 | -104.5885
9.4894 | -104.5619
9.4932 | -104.5429
9.4969 | -104.5306
9.5007 | -104.523
9.5044 | -104.5182
9.5082 | -104.5153
9.5119 | -104.5151
9.5157 | -104.515
9.5194 | -104.5223
9.5232 | -104.5274
9.5269 | -104.5314
9.5307 | -104.5363
9.5344 | -104.5426
9.5382 | -104.5432
9.5419 | -104.5473
9.5457 | -104.5497
9.5494 | -104.5494
9.5532 | -104.5523
9.5569 | -104.5533
9.5607 | -104.5516
9.5644 | -104.552
9.5682 | -104.5513
9.5719 | -104.5519
9.5757 | -104.5537
9.5794 | -104.5575
9.5832 | -104.5646
9.5869 | -104.5764
9.5907 | -104.5881
9.5944 | -104.6023
9.5982 | -104.6255
9.602 | -104.6464
9.6057 | -104.6665
9.6095 | -104.6881
9.6132 | -104.717
9.617 | -104.7444
9.6207 | -104.7744
9.6245 | -104.8061
9.6282 | -104.8381
9.632 | -104.8665
9.6357 | -104.8896
9.6395 | -104.9092
9.6432 | -104.9271
9.647 | -104.9457
9.6507 | -104.96
9.6545 | -104.9729
9.6582 | -104.9837
9.662 | -104.9929
9.6657 | -104.9995
9.6695 | -105.0045
9.6732 | -105.0055
9.677 | -105.0063
9.6807 | -104.9994
9.6845 | -104.9936
9.6882 | -104.988
9.692 | -104.9778
9.6957 | -104.9697
9.6995 | -104.9663
9.7032 | -104.9618
9.707 | -104.961
9.7107 | -104.9618
9.7145 | -104.9663
9.7182 | -104.9662
9.722 | -104.9668
9.7257 | -104.969
9.7295 | -104.9714
9.7332 | -104.9721
9.737 | -104.9708
9.7407 | -104.9723
9.7445 | -104.9732
9.7482 | -104.972
9.752 | -104.973
9.7557 | -104.9688
9.7595 | -104.9716
9.7632 | -104.9763
9.767 | -104.9808
9.7707 | -104.9872
9.7745 | -104.992
9.7782 | -105.003
9.782 | -105.0169
9.7857 | -105.0304
9.7895 | -105.0456
9.7932 | -105.0622
9.797 | -105.0803
9.8007 | -105.1006
9.8045 | -105.1211
9.8082 | -105.1437
9.812 | -105.1711
9.8157 | -105.1999
9.8195 | -105.2246
9.8232 | -105.2512
9.827 | -105.2769
9.8307 | -105.298
9.8345 | -105.321
9.8382 | -105.3358
9.842 | -105.3433
9.8458 | -105.3484
9.8495 | -105.3461
9.8533 | -105.335
9.857 | -105.3181
9.8608 | -105.2955
9.8645 | -105.2647
9.8683 | -105.2253
9.872 | -105.1806
9.8758 | -105.1332
9.8795 | -105.0828
9.8833 | -105.0332
9.887 | -104.9769
9.8908 | -104.9206
9.8945 | -104.8677
9.8983 | -104.8135
9.902 | -104.7651
9.9058 | -104.7175
9.9095 | -104.6758
9.9133 | -104.6441
9.917 | -104.6123
9.9208 | -104.5839
9.9245 | -104.5583
9.9283 | -104.5427
9.932 | -104.5305
9.9358 | -104.5225
9.9395 | -104.5216
9.9433 | -104.5207
9.947 | -104.5194
9.9508 | -104.5222
9.9545 | -104.5256
9.9583 | -104.5249
9.962 | -104.5269
9.9658 | -104.5329
9.9695 | -104.5374
9.9733 | -104.5388
9.977 | -104.5416
9.9808 | -104.5487
9.9845 | -104.559
9.9883 | -104.5681
9.992 | -104.5784
9.9958 | -104.591
9.9995 | -104.6084
10.0033 | -104.6254
10.007 | -104.6445
10.0108 | -104.6627
10.0145 | -104.6853
10.0183 | -104.706
10.022 | -104.7295
10.0258 | -104.7528
10.0295 | -104.7774
10.0333 | -104.8052
10.037 | -104.8311
10.0408 | -104.8583
10.0445 | -104.8829
10.0483 | -104.9042
10.052 | -104.9246
10.0558 | -104.9429
10.0595 | -104.9607
10.0633 | -104.9799
10.067 | -105.0016
10.0708 | -105.0212
10.0745 | -105.0337
10.0783 | -105.0493
10.082 | -105.063
10.0858 | -105.0719
10.0895 | -105.0848
10.0933 | -105.1007
10.0971 | -105.1139
10.1008 | -105.1211
10.1046 | -105.1299
10.1083 | -105.1377
10.1121 | -105.1454
10.1158 | -105.1539
10.1196 | -105.1603
10.1233 | -105.1622
10.1271 | -105.1649
10.1308 | -105.1688
10.1346 | -105.1773
10.1383 | -105.1876
10.1421 | -105.196
10.1458 | -105.1997
10.1496 | -105.2067
10.1533 | -105.2187
10.1571 | -105.2312
10.1608 | -105.2381
10.1646 | -105.2442
10.1683 | -105.245
10.1721 | -105.2437
10.1758 | -105.2394
10.1796 | -105.2355
10.1833 | -105.232
10.1871 | -105.221
10.1908 | -105.2118
10.1946 | -105.2035
10.1983 | -105.1935
10.2021 | -105.1846
10.2058 | -105.1739
10.2096 | -105.1619
10.2133 | -105.1496
10.2171 | -105.1424
10.2208 | -105.14
10.2246 | -105.1383
10.2283 | -105.1387
10.2321 | -105.1428
10.2358 | -105.1462
10.2396 | -105.1503
10.2433 | -105.156
10.2471 | -105.1687
10.2508 | -105.1806
10.2546 | -105.1879
10.2583 | -105.1947
10.2621 | -105.2036
10.2658 | -105.2118
10.2696 | -105.2217
10.2733 | -105.2259
10.2771 | -105.2311
10.2808 | -105.234
10.2846 | -105.2372
10.2883 | -105.239
10.2921 | -105.2428
10.2958 | -105.2495
10.2996 | -105.2548
10.3033 | -105.2537
10.3071 | -105.2511
10.3108 | -105.2528
10.3146 | -105.2569
10.3183 | -105.2568
10.3221 | -105.2548
10.3258 | -105.252
10.3296 | -105.2477
10.3333 | -105.2442
10.3371 | -105.2421
10.3408 | -105.2368
10.3446 | -105.2306
10.3484 | -105.2245
10.3521 | -105.2179
10.3559 | -105.2133
10.3596 | -105.216
10.3634 | -105.2154
10.3671 | -105.2127
10.3709 | -105.2104
10.3746 | -105.2096
10.3784 | -105.2122
10.3821 | -105.2138
10.3859 | -105.2184
10.3896 | -105.2237
10.3934 | -105.228
10.3971 | -105.2298
10.4009 | -105.2342
10.4046 | -105.2402
10.4084 | -105.2436
10.4121 | -105.2464
10.4159 | -105.2516
10.4196 | -105.2566
10.4234 | -105.2601
10.4271 | -105.2579
10.4309 | -105.2513
10.4346 | -105.2469
10.4384 | -105.2394
10.4421 | -105.2319
10.4459 | -105.2223
10.4496 | -105.2095
10.4534 | -105.1928
10.4571 | -105.1768
10.4609 | -105.1604
10.4646 | -105.1431
10.4684 | -105.1252
10.4721 | -105.1113
10.4759 | -105.1011
10.4796 | -105.0917
10.4834 | -105.0876
10.4871 | -105.0841
10.4909 | -105.0849
10.4946 | -105.0958
10.4984 | -105.1096
10.5021 | -105.1283
10.5059 | -105.1494
10.5096 | -105.1755
10.5134 | -105.1999
10.5171 | -105.2283
10.5209 | -105.2608
10.5246 | -105.2976
10.5284 | -105.3333
10.5321 | -105.3657
10.5359 | -105.3996
10.5396 | -105.427
10.5434 | -105.4494
10.5471 | -105.4683
10.5509 | -105.4802
10.5546 | -105.4845
10.5584 | -105.4838
10.5621 | -105.4785
10.5659 | -105.4653
10.5696 | -105.4453
10.5734 | -105.4198
10.5771 | -105.3899
10.5809 | -105.363
10.5846 | -105.3353
10.5884 | -105.2995
10.5922 | -105.266
10.5959 | -105.2336
10.5997 | -105.2038
10.6034 | -105.1804
10.6072 | -105.1633
10.6109 | -105.1489
10.6147 | -105.1418
10.6184 | -105.1398
10.6222 | -105.1446
10.6259 | -105.1582
10.6297 | -105.1776
10.6334 | -105.204
10.6372 | -105.2336
10.6409 | -105.2721
10.6447 | -105.3139
10.6484 | -105.3562
10.6522 | -105.4018
10.6559 | -105.4488
10.6597 | -105.498
10.6634 | -105.543
10.6672 | -105.5891
10.6709 | -105.6369
10.6747 | -105.6787
10.6784 | -105.7096
10.6822 | -105.74
10.6859 | -105.768
10.6897 | -105.7919
10.6934 | -105.8109
10.6972 | -105.8256
10.7009 | -105.8354
10.7047 | -105.8366
10.7084 | -105.8323
10.7122 | -105.8263
10.7159 | -105.8142
10.7197 | -105.7963
10.7234 | -105.7734
10.7272 | -105.7418
10.7309 | -105.707
10.7347 | -105.6729
10.7384 | -105.6398
10.7422 | -105.6065
10.7459 | -105.5686
10.7497 | -105.531
10.7534 | -105.4952
10.7572 | -105.4597
10.7609 | -105.428
10.7647 | -105.399
10.7684 | -105.3722
10.7722 | -105.3471
10.7759 | -105.3256
10.7797 | -105.3034
10.7834 | -105.2892
10.7872 | -105.2779
10.7909 | -105.2719
10.7947 | -105.2693
10.7984 | -105.2732
10.8022 | -105.2807
10.8059 | -105.2928
10.8097 | -105.3105
10.8134 | -105.3285
10.8172 | -105.351
10.8209 | -105.3756
10.8247 | -105.3961
10.8285 | -105.4155
10.8322 | -105.4356
10.836 | -105.4557
10.8397 | -105.4719
10.8435 | -105.4879
10.8472 | -105.5002
10.851 | -105.511
10.8547 | -105.5192
10.8585 | -105.5206
10.8622 | -105.5215
10.866 | -105.5204
10.8697 | -105.5095
10.8735 | -105.4958
10.8772 | -105.4827
10.881 | -105.4712
10.8847 | -105.4595
10.8885 | -105.4455
10.8922 | -105.424
10.896 | -105.3991
10.8997 | -105.3764
10.9035 | -105.357
10.9072 | -105.3385
10.911 | -105.3202
10.9147 | -105.3051
10.9185 | -105.2924
10.9222 | -105.285
10.926 | -105.2791
10.9297 | -105.2803
10.9335 | -105.2842
10.9372 | -105.2907
10.941 | -105.2959
10.9447 | -105.3081
10.9485 | -105.3237
10.9522 | -105.3417
10.956 | -105.3627
10.9597 | -105.3835
10.9635 | -105.4064
10.9672 | -105.4293
10.971 | -105.4538
10.9747 | -105.4763
10.9785 | -105.4999
10.9822 | -105.5208
10.986 | -105.5449
10.9897 | -105.5654
10.9935 | -105.5858
10.9972 | -105.6053
11.001 | -105.6262
11.0047 | -105.644
11.0085 | -105.6594
11.0122 | -105.6755
11.016 | -105.687
11.0197 | -105.7058
11.0235 | -105.728
11.0272 | -105.7468
11.031 | -105.7679
11.0347 | -105.7905
11.0385 | -105.8138
11.0422 | -105.8352
11.046 | -105.8592
11.0497 | -105.8799
11.0535 | -105.9056
11.0572 | -105.9298
11.061 | -105.9518
11.0647 | -105.9705
11.0685 | -105.9858
11.0722 | -105.9999
11.076 | -106.0035
11.0798 | -106.0047
11.0835 | -106.0053
11.0873 | -106.0046
11.091 | -106.0004
11.0948 | -105.9928
11.0985 | -105.9787
11.1023 | -105.9646
11.106 | -105.9476
11.1098 | -105.9275
11.1135 | -105.9026
11.1173 | -105.8745
11.121 | -105.8454
11.1248 | -105.8165
11.1285 | -105.7879
11.1323 | -105.7596
11.136 | -105.7322
11.1398 | -105.7025
11.1435 | -105.6707
11.1473 | -105.64
11.151 | -105.6118
11.1548 | -105.5867
11.1585 | -105.5678
11.1623 | -105.549
11.166 | -105.5315
11.1698 | -105.52
11.1735 | -105.5102
11.1773 | -105.5074
11.181 | -105.5036
11.1848 | -105.5008
11.1885 | -105.5027
11.1923 | -105.5126
11.196 | -105.5249
11.1998 | -105.5375
11.2035 | -105.5522
11.2073 | -105.5685
11.211 | -105.5901
11.2148 | -105.6116
11.2185 | -105.6317
11.2223 | -105.6508
11.226 | -105.6693
11.2298 | -105.689
11.2335 | -105.7095
11.2373 | -105.7301
11.241 | -105.7456
11.2448 | -105.7602
11.2485 | -105.771
11.2523 | -105.7787
11.256 | -105.7875
11.2598 | -105.8007
11.2635 | -105.8055
11.2673 | -105.8044
11.271 | -105.804
11.2748 | -105.7998
11.2785 | -105.7945
11.2823 | -105.7887
11.286 | -105.7796
11.2898 | -105.7684
11.2935 | -105.7587
11.2973 | -105.7492
11.301 | -105.742
11.3048 | -105.7355
11.3085 | -105.7266
11.3123 | -105.7181
11.316 | -105.7103
11.3198 | -105.7006
11.3235 | -105.6929
11.3273 | -105.6856
11.3311 | -105.6772
11.3348 | -105.6692
11.3386 | -105.6617
11.3423 | -105.6532
11.3461 | -105.6466
11.3498 | -105.6402
11.3536 | -105.6315
11.3573 | -105.6234
11.3611 | -105.6198
11.3648 | -105.6155
11.3686 | -105.6097
11.3723 | -105.6017
11.3761 | -105.5953
11.3798 | -105.5881
11.3836 | -105.5844
11.3873 | -105.5816
11.3911 | -105.5777
11.3948 | -105.5688
11.3986 | -105.5575
11.4023 | -105.5524
11.4061 | -105.547
11.4098 | -105.5381
11.4136 | -105.5323
11.4173 | -105.5261
11.4211 | -105.52
11.4248 | -105.5119
11.4286 | -105.5058
11.4323 | -105.504
11.4361 | -105.5007
11.4398 | -105.5025
11.4436 | -105.5016
11.4473 | -105.5027
11.4511 | -105.5039
11.4548 | -105.5041
11.4586 | -105.5034
11.4623 | -105.5034
11.4661 | -105.5066
11.4698 | -105.5107
11.4736 | -105.5166
11.4773 | -105.5257
11.4811 | -105.5335
11.4848 | -105.5429
11.4886 | -105.5534
11.4923 | -105.5661
11.4961 | -105.5794
11.4998 | -105.5913
11.5036 | -105.6009
11.5073 | -105.6104
11.5111 | -105.6197
11.5148 | -105.6284
11.5186 | -105.6332
11.5223 | -105.635
11.5261 | -105.6379
11.5298 | -105.6394
11.5336 | -105.6408
11.5373 | -105.6458
11.5411 | -105.649
11.5448 | -105.6452
11.5486 | -105.6369
11.5523 | -105.6314
11.5561 | -105.6277
11.5598 | -105.6197
11.5636 | -105.6097
11.5673 | -105.5989
11.5711 | -105.5846
11.5749 | -105.5736
11.5786 | -105.564
11.5824 | -105.5508
11.5861 | -105.5396
11.5899 | -105.528
11.5936 | -105.5173
11.5974 | -105.5085
11.6011 | -105.4976
11.6049 | -105.488
11.6086 | -105.4786
11.6124 | -105.4684
11.6161 | -105.459
11.6199 | -105.4513
11.6236 | -105.4477
11.6274 | -105.4435
11.6311 | -105.4419
11.6349 | -105.4399
11.6386 | -105.4377
11.6424 | -105.4397
11.6461 | -105.4409
11.6499 | -105.4431
11.6536 | -105.4496
11.6574 | -105.4562
11.6611 | -105.4633
11.6649 | -105.475
11.6686 | -105.4892
11.6724 | -105.5041
11.6761 | -105.5218
11.6799 | -105.5395
11.6836 | -105.558
11.6874 | -105.5781
11.6911 | -105.5986
11.6949 | -105.616
11.6986 | -105.6307
11.7024 | -105.6442
11.7061 | -105.6585
11.7099 | -105.6721
11.7136 | -105.6793
11.7174 | -105.6869
11.7211 | -105.6883
11.7249 | -105.6805
11.7286 | -105.6776
11.7324 | -105.675
11.7361 | -105.67
11.7399 | -105.6625
11.7436 | -105.6498
11.7474 | -105.6367
11.7511 | -105.6262
11.7549 | -105.6149
11.7586 | -105.6031
11.7624 | -105.594
11.7661 | -105.5838
11.7699 | -105.5705
11.7736 | -105.5594
11.7774 | -105.5532
11.7811 | -105.5513
11.7849 | -105.5499
11.7886 | -105.5494
11.7924 | -105.5495
11.7961 | -105.5493
11.7999 | -105.5478
11.8036 | -105.5485
11.8074 | -105.5519
11.8111 | -105.5524
11.8149 | -105.5559
11.8187 | -105.5623
11.8224 | -105.5646
11.8262 | -105.5677
11.8299 | -105.5739
11.8337 | -105.5746
11.8374 | -105.5777
11.8412 | -105.5798
11.8449 | -105.5792
11.8487 | -105.5792
11.8524 | -105.5787
11.8562 | -105.5794
11.8599 | -105.5817
11.8637 | -105.5824
11.8674 | -105.5823
11.8712 | -105.5854
11.8749 | -105.5876
11.8787 | -105.5923
11.8824 | -105.6032
11.8862 | -105.6129
11.8899 | -105.6216
11.8937 | -105.6346
11.8974 | -105.6438
11.9012 | -105.6597
11.9049 | -105.6704
11.9087 | -105.6835
11.9124 | -105.6989
11.9162 | -105.7123
11.9199 | -105.7256
11.9237 | -105.7398
11.9274 | -105.7568
11.9312 | -105.7766
11.9349 | -105.7956
11.9387 | -105.8171
11.9424 | -105.8406
11.9462 | -105.8577
11.9499 | -105.8696
11.9537 | -105.8837
11.9574 | -105.8958
11.9612 | -105.9065
11.9649 | -105.9115
11.9687 | -105.9175
11.9724 | -105.9234
11.9762 | -105.9222
11.9799 | -105.9168
11.9837 | -105.9079
11.9874 | -105.8948
11.9912 | -105.8767
11.9949 | -105.858
11.9987 | -105.8389
12.0024 | -105.8158
12.0062 | -105.792
12.0099 | -105.7682
12.0137 | -105.7494
12.0174 | -105.7315
12.0212 | -105.7122
12.0249 | -105.6972
12.0287 | -105.6836
12.0324 | -105.6691
12.0362 | -105.6639
12.0399 | -105.6652
12.0437 | -105.6666
12.0474 | -105.6675
12.0512 | -105.6719
12.0549 | -105.6819
12.0587 | -105.6951
12.0625 | -105.7117
12.0662 | -105.7296
12.07 | -105.7484
12.0737 | -105.7674
12.0775 | -105.7878
12.0812 | -105.8089
12.085 | -105.8247
12.0887 | -105.8381
12.0925 | -105.8542
12.0962 | -105.8666
12.1 | -105.8756
12.1037 | -105.8823
12.1075 | -105.8831
12.1112 | -105.8805
12.115 | -105.874
12.1187 | -105.868
12.1225 | -105.8642
12.1262 | -105.8603
12.13 | -105.8506
12.1337 | -105.8374
12.1375 | -105.8253
12.1412 | -105.8151
12.145 | -105.8012
12.1487 | -105.7815
12.1525 | -105.766
12.1562 | -105.7496
12.16 | -105.735
12.1637 | -105.7199
12.1675 | -105.7077
12.1712 | -105.6958
12.175 | -105.6849
12.1787 | -105.678
12.1825 | -105.6775
12.1862 | -105.6763
12.19 | -105.6761
12.1937 | -105.6804
12.1975 | -105.6887
12.2012 | -105.6986
12.205 | -105.7097
12.2087 | -105.7272
12.2125 | -105.7426
12.2162 | -105.7587
12.22 | -105.7802
12.2237 | -105.8075
12.2275 | -105.836
12.2312 | -105.8617
12.235 | -105.8856
12.2387 | -105.9102
12.2425 | -105.9355
12.2462 | -105.9636
12.25 | -105.9937
12.2537 | -106.0212
12.2575 | -106.0469
12.2612 | -106.0749
12.265 | -106.1027
12.2687 | -106.1278
12.2725 | -106.1515
12.2762 | -106.1666
12.28 | -106.1802
12.2837 | -106.1981
12.2875 | -106.2132
12.2912 | -106.2257
12.295 | -106.2323
12.2987 | -106.2391
12.3025 | -106.243
12.3063 | -106.2423
12.31 | -106.2445
12.3138 | -106.2436
12.3175 | -106.2393
12.3213 | -106.2344
12.325 | -106.2292
12.3288 | -106.2207
12.3325 | -106.2135
12.3363 | -106.2045
12.34 | -106.1978
12.3438 | -106.1901
12.3475 | -106.1795
12.3513 | -106.1673
12.355 | -106.1567
12.3588 | -106.1446
12.3625 | -106.1355
12.3663 | -106.1277
12.37 | -106.1189
12.3738 | -106.1072
12.3775 | -106.0984
12.3813 | -106.0918
12.385 | -106.0826
12.3888 | -106.0747
12.3925 | -106.0691
12.3963 | -106.0607
12.4 | -106.0523
12.4038 | -106.0452
12.4075 | -106.0348
12.4113 | -106.0253
12.415 | -106.0162
12.4188 | -106.0046
12.4225 | -105.9897
12.4263 | -105.9706
12.43 | -105.9511
12.4338 | -105.9358
12.4375 | -105.9196
12.4413 | -105.9029
12.445 | -105.8893
12.4488 | -105.8754
12.4525 | -105.865
12.4563 | -105.859
12.46 | -105.8521
12.4638 | -105.8401
12.4675 | -105.8337
12.4713 | -105.8312
12.475 | -105.8302
12.4788 | -105.8283
12.4825 | -105.8282
12.4863 | -105.8331
12.49 | -105.8394
12.4938 | -105.8453
12.4975 | -105.8526
12.5013 | -105.8606
12.505 | -105.871
12.5088 | -105.8819
12.5125 | -105.896
12.5163 | -105.9086
12.52 | -105.9178
12.5238 | -105.929
12.5275 | -105.9436
12.5313 | -105.9565
12.535 | -105.9674
12.5388 | -105.9767
12.5425 | -105.9827
12.5463 | -105.9901
12.55 | -105.9991
12.5538 | -106.0078
12.5576 | -106.0132
12.5613 | -106.0135
12.5651 | -106.014
12.5688 | -106.0135
12.5726 | -106.0075
12.5763 | -106.0004
12.5801 | -105.9958
12.5838 | -105.9906
12.5876 | -105.9868
12.5913 | -105.9812
12.5951 | -105.9747
12.5988 | -105.9681
12.6026 | -105.959
12.6063 | -105.954
12.6101 | -105.9512
12.6138 | -105.9486
12.6176 | -105.9463
12.6213 | -105.9453
12.6251 | -105.9461
12.6288 | -105.9489
12.6326 | -105.9511
12.6363 | -105.955
12.6401 | -105.9635
12.6438 | -105.9708
12.6476 | -105.981
12.6513 | -105.9908
12.6551 | -106.0021
12.6588 | -106.0144
12.6626 | -106.0275
12.6663 | -106.0414
12.6701 | -106.0548
12.6738 | -106.0686
12.6776 | -106.0829
12.6813 | -106.0983
12.6851 | -106.1137
12.6888 | -106.1244
12.6926 | -106.1368
12.6963 | -106.1447
12.7001 | -106.153
12.7038 | -106.1625
12.7076 | -106.1663
12.7113 | -106.1724
12.7151 | -106.1782
12.7188 | -106.1802
12.7226 | -106.1807
12.7263 | -106.1823
12.7301 | -106.1752
12.7338 | -106.1652
12.7376 | -106.1584
12.7413 | -106.1514
12.7451 | -106.1425
12.7488 | -106.1308
12.7526 | -106.1147
12.7563 | -106.1017
12.7601 | -106.0897
12.7638 | -106.073
12.7676 | -106.0553
12.7713 | -106.0402
12.7751 | -106.027
12.7788 | -106.0143
12.7826 | -106.0027
12.7863 | -105.9937
12.7901 | -105.9812
12.7938 | -105.9676
12.7976 | -105.9544
12.8013 | -105.9456
12.8051 | -105.9356
12.8089 | -105.9261
12.8126 | -105.9183
12.8164 | -105.9088
12.8201 | -105.899
12.8239 | -105.893
12.8276 | -105.8895
12.8314 | -105.8829
12.8351 | -105.8736
12.8389 | -105.8659
12.8426 | -105.8598
12.8464 | -105.8521
12.8501 | -105.8453
12.8539 | -105.84
12.8576 | -105.8294
12.8614 | -105.8206
12.8651 | -105.8131
12.8689 | -105.8048
12.8726 | -105.802
12.8764 | -105.7954
12.8801 | -105.7917
12.8839 | -105.7897
12.8876 | -105.7895
12.8914 | -105.7912
12.8951 | -105.7908
12.8989 | -105.7933
12.9026 | -105.7973
12.9064 | -105.801
12.9101 | -105.7984
12.9139 | -105.8015
12.9176 | -105.809
12.9214 | -105.8105
12.9251 | -105.8137
12.9289 | -105.8173
12.9326 | -105.8174
12.9364 | -105.8179
12.9401 | -105.8188
12.9439 | -105.8193
12.9476 | -105.8205
12.9514 | -105.8206
12.9551 | -105.8199
12.9589 | -105.8181
12.9626 | -105.814
12.9664 | -105.8113
12.9701 | -105.808
12.9739 | -105.8039
12.9776 | -105.8001
12.9814 | -105.7947
12.9851 | -105.7902
12.9889 | -105.7826
12.9926 | -105.7732
12.9964 | -105.7643
13.0001 | -105.757
13.0039 | -105.7501
13.0076 | -105.7428
13.0114 | -105.7343
13.0151 | -105.7255
13.0189 | -105.7187
13.0226 | -105.7141
13.0264 | -105.707
13.0301 | -105.6996
13.0339 | -105.6969
13.0376 | -105.6931
13.0414 | -105.6894
13.0451 | -105.6891
13.0489 | -105.6879
13.0527 | -105.6843
13.0564 | -105.6818
13.0602 | -105.6798
13.0639 | -105.6804
13.0677 | -105.6814
13.0714 | -105.6849
13.0752 | -105.6897
13.0789 | -105.6976
13.0827 | -105.7065
13.0864 | -105.7153
13.0902 | -105.7253
13.0939 | -105.739
13.0977 | -105.7579
13.1014 | -105.777
13.1052 | -105.7941
13.1089 | -105.8116
13.1127 | -105.8302
13.1164 | -105.8548
13.1202 | -105.8805
13.1239 | -105.9065
13.1277 | -105.9297
13.1314 | -105.9516
13.1352 | -105.9719
13.1389 | -105.9918
13.1427 | -106.0089
13.1464 | -106.024
13.1502 | -106.0405
13.1539 | -106.0558
13.1577 | -106.0669
13.1614 | -106.0773
13.1652 | -106.0867
13.1689 | -106.0943
13.1727 | -106.0986
13.1764 | -106.1043
13.1802 | -106.1085
13.1839 | -106.1098
13.1877 | -106.1093
13.1914 | -106.1053
13.1952 | -106.1042
13.1989 | -106.0986
13.2027 | -106.0897
13.2064 | -106.0848
13.2102 | -106.0803
13.2139 | -106.0754
13.2177 | -106.069
13.2214 | -106.0619
13.2252 | -106.05
13.2289 | -106.0393
13.2327 | -106.0287
13.2364 | -106.0179
13.2402 | -106.0059
13.2439 | -105.9975
13.2477 | -105.9904
13.2514 | -105.9797
13.2552 | -105.9685
13.2589 | -105.9583
13.2627 | -105.9488
13.2664 | -105.9397
13.2702 | -105.932
13.2739 | -105.9279
13.2777 | -105.9221
13.2814 | -105.9202
13.2852 | -105.9207
13.2889 | -105.923
13.2927 | -105.9231
13.2965 | -105.9238
13.3002 | -105.9282
13.304 | -105.9301
13.3077 | -105.9264
13.3115 | -105.9276
13.3152 | -105.9288
13.319 | -105.9262
13.3227 | -105.9225
13.3265 | -105.9161
13.3302 | -105.9076
13.334 | -105.9008
13.3377 | -105.8932
13.3415 | -105.8807
13.3452 | -105.8648
13.349 | -105.8482
13.3527 | -105.8301
13.3565 | -105.8085
13.3602 | -105.7876
13.364 | -105.7648
13.3677 | -105.7416
13.3715 | -105.7175
13.3752 | -105.6963
13.379 | -105.6741
13.3827 | -105.6514
13.3865 | -105.6295
13.3902 | -105.6089
13.394 | -105.591
13.3977 | -105.5747
13.4015 | -105.5607
13.4052 | -105.5467
13.409 | -105.5349
13.4127 | -105.523
13.4165 | -105.5172
13.4202 | -105.5146
13.424 | -105.5144
13.4277 | -105.514
13.4315 | -105.5169
13.4352 | -105.5195
13.439 | -105.5224
13.4427 | -105.528
13.4465 | -105.5367
13.4502 | -105.5462
13.454 | -105.5573
13.4577 | -105.5646
13.4615 | -105.5725
13.4652 | -105.5804
13.469 | -105.5922
13.4727 | -105.6009
13.4765 | -105.6064
13.4802 | -105.6116
13.484 | -105.6199
13.4877 | -105.6269
13.4915 | -105.6344
13.4952 | -105.6417
13.499 | -105.6505
13.5027 | -105.6549
13.5065 | -105.6604
13.5102 | -105.6643
13.514 | -105.6692
13.5177 | -105.6752
13.5215 | -105.6815
13.5252 | -105.6879
13.529 | -105.6919
13.5327 | -105.6953
13.5365 | -105.6981
13.5403 | -105.701
13.544 | -105.7069
13.5478 | -105.7157
13.5515 | -105.7257
13.5553 | -105.7357
13.559 | -105.7462
13.5628 | -105.7536
13.5665 | -105.763
13.5703 | -105.7719
13.574 | -105.7797
13.5778 | -105.7845
13.5815 | -105.79
13.5853 | -105.7973
13.589 | -105.7981
13.5928 | -105.7954
13.5965 | -105.7955
13.6003 | -105.7937
13.604 | -105.794
13.6078 | -105.794
13.6115 | -105.7858
13.6153 | -105.7767
13.619 | -105.7677
13.6228 | -105.7587
13.6265 | -105.7474
13.6303 | -105.7389
13.634 | -105.7291
13.6378 | -105.7195
13.6415 | -105.7082
13.6453 | -105.6962
13.649 | -105.6831
13.6528 | -105.6761
13.6565 | -105.6664
13.6603 | -105.6571
13.664 | -105.646
13.6678 | -105.6348
13.6715 | -105.6254
13.6753 | -105.6154
13.679 | -105.6065
13.6828 | -105.5989
13.6865 | -105.5931
13.6903 | -105.5863
13.694 | -105.5814
13.6978 | -105.5752
13.7015 | -105.5632
13.7053 | -105.5591
13.709 | -105.5543
13.7128 | -105.5512
13.7165 | -105.5484
13.7203 | -105.546
13.724 | -105.5442
13.7278 | -105.5411
13.7315 | -105.5395
13.7353 | -105.5396
13.739 | -105.537
13.7428 | -105.5312
13.7465 | -105.5244
13.7503 | -105.5178
13.754 | -105.5113
13.7578 | -105.5087
13.7615 | -105.5061
13.7653 | -105.5033
13.769 | -105.4996
13.7728 | -105.4971
13.7765 | -105.4987
13.7803 | -105.4996
13.7841 | -105.4995
13.7878 | -105.5031
13.7916 | -105.5077
13.7953 | -105.5115
13.7991 | -105.5165
13.8028 | -105.5217
13.8066 | -105.5298
13.8103 | -105.5427
13.8141 | -105.5539
13.8178 | -105.5653
13.8216 | -105.5787
13.8253 | -105.5933
13.8291 | -105.606
13.8328 | -105.6165
13.8366 | -105.6282
13.8403 | -105.6415
13.8441 | -105.6559
13.8478 | -105.6681
13.8516 | -105.6792
13.8553 | -105.6926
13.8591 | -105.7038
13.8628 | -105.7138
13.8666 | -105.7255
13.8703 | -105.7369
13.8741 | -105.745
13.8778 | -105.7494
13.8816 | -105.7559
13.8853 | -105.7631
13.8891 | -105.7675
13.8928 | -105.7716
13.8966 | -105.7767
13.9003 | -105.785
13.9041 | -105.7882
13.9078 | -105.7895
13.9116 | -105.7908
13.9153 | -105.7934
13.9191 | -105.7973
13.9228 | -105.7984
13.9266 | -105.8002
13.9303 | -105.803
13.9341 | -105.8044
13.9378 | -105.8068
13.9416 | -105.8086
13.9453 | -105.8048
13.9491 | -105.8014
13.9528 | -105.7983
13.9566 | -105.7911
13.9603 | -105.7676
13.9641 | -105.7292
13.9678 | -105.6965
13.9716 | -105.67
13.9753 | -105.6396
13.9791 | -105.6062
13.9828 | -105.5693
13.9866 | -105.5295
13.9903 | -105.4855
13.9941 | -105.4382
13.9978 | -105.3893
14.0016 | -105.3336
14.0053 | -105.2772
14.0091 | -105.22
14.0128 | -105.1625
14.0166 | -105.107
14.0203 | -105.0505
14.0241 | -104.9935
14.0278 | -104.9407
14.0316 | -104.8954
14.0354 | -104.8522
14.0391 | -104.8131
14.0429 | -104.7784
14.0466 | -104.7495
14.0504 | -104.7282
14.0541 | -104.7128
14.0579 | -104.7059
14.0616 | -104.7095
14.0654 | -104.7522
14.0691 | -104.7661
14.0729 | -104.789
14.0766 | -104.821
14.0804 | -104.8589
14.0841 | -104.9012
14.0879 | -104.9475
14.0916 | -104.9997
14.0954 | -105.0552
14.0991 | -105.1129
14.1029 | -105.1733
14.1066 | -105.232
14.1104 | -105.2943
14.1141 | -105.3558
14.1179 | -105.4127
14.1216 | -105.4679
14.1254 | -105.5262
14.1291 | -105.5773
14.1329 | -105.6232
14.1366 | -105.6688
14.1404 | -105.7115
14.1441 | -105.7497
14.1479 | -105.7855
14.1516 | -105.817
14.1554 | -105.8425
14.1591 | -105.8653
14.1629 | -105.9
14.1666 | -105.936
14.1704 | -105.9564
14.1741 | -105.9612
14.1779 | -105.965
14.1816 | -105.968
14.1854 | -105.9718
14.1891 | -105.9768
14.1929 | -105.9793
14.1966 | -105.9811
14.2004 | -105.9818
14.2041 | -105.9852
14.2079 | -105.9843
14.2116 | -105.9798
14.2154 | -105.9769
14.2191 | -105.9712
14.2229 | -105.9669
14.2266 | -105.9639
14.2304 | -105.9586
14.2341 | -105.9534
14.2379 | -105.9503
14.2416 | -105.9469
14.2454 | -105.9422
14.2491 | -105.9376
14.2529 | -105.9324
14.2566 | -105.9269
14.2604 | -105.9183
14.2641 | -105.9091
14.2679 | -105.9004
14.2716 | -105.8882
14.2754 | -105.8751
14.2791 | -105.8601
14.2829 | -105.8441
14.2867 | -105.8294
14.2904 | -105.816
14.2942 | -105.8023
14.2979 | -105.7896
14.3017 | -105.7747
14.3054 | -105.7628
14.3092 | -105.7496
14.3129 | -105.7394
14.3167 | -105.7309
14.3204 | -105.7229
14.3242 | -105.7205
14.3279 | -105.7181
14.3317 | -105.7146
14.3354 | -105.7161
14.3392 | -105.7196
14.3429 | -105.7237
14.3467 | -105.7276
14.3504 | -105.7332
14.3542 | -105.7412
14.3579 | -105.7462
14.3617 | -105.7549
14.3654 | -105.7647
14.3692 | -105.7763
14.3729 | -105.7894
14.3767 | -105.8002
14.3804 | -105.8115
14.3842 | -105.8245
14.3879 | -105.8394
14.3917 | -105.8532
14.3954 | -105.8654
14.3992 | -105.8764
14.4029 | -105.8876
14.4067 | -105.9
14.4104 | -105.9121
14.4142 | -105.9244
14.4179 | -105.936
14.4217 | -105.9454
14.4254 | -105.9562
14.4292 | -105.964
14.4329 | -105.9682
14.4367 | -105.97
14.4404 | -105.9686
14.4442 | -105.9674
14.4479 | -105.9656
14.4517 | -105.9644
14.4554 | -105.9657
14.4592 | -105.9652
14.4629 | -105.9616
14.4667 | -105.962
14.4704 | -105.9652
14.4742 | -105.9679
14.4779 | -105.9717
14.4817 | -105.9768
14.4854 | -105.9777
14.4892 | -105.9808
14.4929 | -105.9832
14.4967 | -105.9831
14.5004 | -105.9861
14.5042 | -105.9912
14.5079 | -105.9966
14.5117 | -105.9986
14.5154 | -106.0012
14.5192 | -106.0026
14.5229 | -105.9996
14.5267 | -105.9995
14.5305 | -105.998
14.5342 | -105.9988
14.538 | -106.002
14.5417 | -106.0062
14.5455 | -106.0106
14.5492 | -106.016
14.553 | -106.0243
14.5567 | -106.0318
14.5605 | -106.0383
14.5642 | -106.0453
14.568 | -106.0511
14.5717 | -106.0576
14.5755 | -106.0635
14.5792 | -106.0686
14.583 | -106.0741
14.5867 | -106.0819
14.5905 | -106.0899
14.5942 | -106.0917
14.598 | -106.0943
14.6017 | -106.0969
14.6055 | -106.103
14.6092 | -106.1063
14.613 | -106.1075
14.6167 | -106.1048
14.6205 | -106.1019
14.6242 | -106.1002
14.628 | -106.1005
14.6317 | -106.0996
14.6355 | -106.0972
14.6392 | -106.0952
14.643 | -106.0951
14.6467 | -106.0952
14.6505 | -106.0974
14.6542 | -106.0967
14.658 | -106.0924
14.6617 | -106.0871
14.6655 | -106.0819
14.6692 | -106.0773
14.673 | -106.0768
14.6767 | -106.0745
14.6805 | -106.0698
14.6842 | -106.0659
14.688 | -106.0643
14.6917 | -106.0621
14.6955 | -106.0597
14.6992 | -106.058
14.703 | -106.0546
14.7067 | -106.0502
14.7105 | -106.0441
14.7142 | -106.0393
14.718 | -106.0355
14.7217 | -106.0322
14.7255 | -106.0245
14.7292 | -106.0149
14.733 | -106.009
14.7367 | -105.9983
14.7405 | -105.9896
14.7442 | -105.9786
14.748 | -105.9659
14.7517 | -105.9545
14.7555 | -105.9448
14.7592 | -105.9345
14.763 | -105.9252
14.7667 | -105.9176
14.7705 | -105.9091
14.7743 | -105.9024
14.778 | -105.8922
14.7818 | -105.8816
14.7855 | -105.8713
14.7893 | -105.8603
14.793 | -105.8494
14.7968 | -105.8391
14.8005 | -105.8257
14.8043 | -105.8179
14.808 | -105.8098
14.8118 | -105.7999
14.8155 | -105.7884
14.8193 | -105.7804
14.823 | -105.7733
14.8268 | -105.7691
14.8305 | -105.7682
14.8343 | -105.7676
14.838 | -105.7645
14.8418 | -105.7615
14.8455 | -105.7627
14.8493 | -105.7627
14.853 | -105.7665
14.8568 | -105.7723
14.8605 | -105.7775
14.8643 | -105.7838
14.868 | -105.792
14.8718 | -105.8038
14.8755 | -105.8155
14.8793 | -105.8221
14.883 | -105.8293
14.8868 | -105.8358
14.8905 | -105.8413
14.8943 | -105.8457
14.898 | -105.8506
14.9018 | -105.8531
14.9055 | -105.855
14.9093 | -105.8594
14.913 | -105.863
14.9168 | -105.8671
14.9205 | -105.8715
14.9243 | -105.874
14.928 | -105.8724
14.9318 | -105.8723
14.9355 | -105.8726
14.9393 | -105.8777
14.943 | -105.8849
14.9468 | -105.8903
14.9505 | -105.8981
14.9543 | -105.9027
14.958 | -105.9073
14.9618 | -105.9154
14.9655 | -105.9232
14.9693 | -105.9304
14.973 | -105.9385
14.9768 | -105.9485
14.9805 | -105.9578
14.9843 | -105.9665
14.988 | -105.9779
14.9918 | -105.9937
14.9955 | -106.009
14.9993 | -106.0224
15.003 | -106.0347
15.0068 | -106.0486
15.0105 | -106.0581
15.0143 | -106.065
15.0181 | -106.0718
15.0218 | -106.0768
15.0256 | -106.0807
15.0293 | -106.0822
15.0331 | -106.079
15.0368 | -106.0749
15.0406 | -106.0666
15.0443 | -106.0525
15.0481 | -106.0315
15.0518 | -106.01
15.0556 | -105.9852
15.0593 | -105.955
15.0631 | -105.9279
15.0668 | -105.8991
15.0706 | -105.8652
15.0743 | -105.8296
15.0781 | -105.7934
15.0818 | -105.7586
15.0856 | -105.7223
15.0893 | -105.6873
15.0931 | -105.6558
15.0968 | -105.6223
15.1006 | -105.5873
15.1043 | -105.5538
15.1081 | -105.5243
15.1118 | -105.4993
15.1156 | -105.4783
15.1193 | -105.4591
15.1231 | -105.4396
15.1268 | -105.4204
15.1306 | -105.4048
15.1343 | -105.3922
15.1381 | -105.3838
15.1418 | -105.3755
15.1456 | -105.3685
15.1493 | -105.3771
15.1531 | -105.3769
15.1568 | -105.3797
15.1606 | -105.385
15.1643 | -105.3919
15.1681 | -105.3989
15.1718 | -105.4091
15.1756 | -105.4251
15.1793 | -105.4393
15.1831 | -105.4497
15.1868 | -105.4646
15.1906 | -105.4821
15.1943 | -105.4984
15.1981 | -105.5174
15.2018 | -105.5378
15.2056 | -105.5563
15.2093 | -105.5726
15.2131 | -105.587
15.2168 | -105.6017
15.2206 | -105.6201
15.2243 | -105.6365
15.2281 | -105.6503
15.2318 | -105.6642
15.2356 | -105.6789
15.2393 | -105.692
15.2431 | -105.7036
15.2468 | -105.7147
15.2506 | -105.7222
15.2543 | -105.7272
15.2581 | -105.7342
15.2619 | -105.7425
15.2656 | -105.7503
15.2694 | -105.7555
15.2731 | -105.7591
15.2769 | -105.7596
15.2806 | -105.7577
15.2844 | -105.757
15.2881 | -105.7578
15.2919 | -105.7576
15.2956 | -105.7562
15.2994 | -105.7524
15.3031 | -105.7514
15.3069 | -105.7523
15.3106 | -105.75
15.3144 | -105.7491
15.3181 | -105.7516
15.3219 | -105.754
15.3256 | -105.7536
15.3294 | -105.7509
15.3331 | -105.7458
15.3369 | -105.7431
15.3406 | -105.7437
15.3444 | -105.7448
15.3481 | -105.7469
15.3519 | -105.7476
15.3556 | -105.7442
15.3594 | -105.7466
15.3631 | -105.7447
15.3669 | -105.7402
15.3706 | -105.7343
15.3744 | -105.7266
15.3781 | -105.715
15.3819 | -105.7001
15.3856 | -105.6869
15.3894 | -105.6742
15.3931 | -105.6605
15.3969 | -105.6426
15.4006 | -105.6256
15.4044 | -105.612
15.4081 | -105.5962
15.4119 | -105.58
15.4156 | -105.5649
15.4194 | -105.5481
15.4231 | -105.5325
15.4269 | -105.5161
15.4306 | -105.499
15.4344 | -105.4849
15.4381 | -105.4721
15.4419 | -105.4588
15.4456 | -105.4462
15.4494 | -105.4381
15.4531 | -105.4317
15.4569 | -105.422
15.4606 | -105.4146
15.4644 | -105.411
15.4681 | -105.4059
15.4719 | -105.4034
15.4756 | -105.4041
15.4794 | -105.4052
15.4831 | -105.4086
15.4869 | -105.4145
15.4906 | -105.4213
15.4944 | -105.428
15.4981 | -105.438
15.5019 | -105.4477
15.5056 | -105.4576
15.5094 | -105.4668
15.5132 | -105.4757
15.5169 | -105.4872
15.5207 | -105.4997
15.5244 | -105.5115
15.5282 | -105.5213
15.5319 | -105.5311
15.5357 | -105.5418
15.5394 | -105.5519
15.5432 | -105.5606
15.5469 | -105.5704
15.5507 | -105.5769
15.5544 | -105.5832
15.5582 | -105.5875
15.5619 | -105.5852
15.5657 | -105.5842
15.5694 | -105.582
15.5732 | -105.5778
15.5769 | -105.5727
15.5807 | -105.5636
15.5844 | -105.5581
15.5882 | -105.5477
15.5919 | -105.5373
15.5957 | -105.5261
15.5994 | -105.5129
15.6032 | -105.5003
15.6069 | -105.486
15.6107 | -105.4719
15.6144 | -105.4587
15.6182 | -105.445
15.6219 | -105.4278
15.6257 | -105.4142
15.6294 | -105.3981
15.6332 | -105.3839
15.6369 | -105.3722
15.6407 | -105.3581
15.6444 | -105.342
15.6482 | -105.3249
15.6519 | -105.3105
15.6557 | -105.2995
15.6594 | -105.2882
15.6632 | -105.2724
15.6669 | -105.2543
15.6707 | -105.2392
15.6744 | -105.225
15.6782 | -105.2106
15.6819 | -105.1951
15.6857 | -105.1794
15.6894 | -105.1577
15.6932 | -105.1423
15.6969 | -105.1259
15.7007 | -105.109
15.7044 | -105.0929
15.7082 | -105.0745
15.7119 | -105.0572
15.7157 | -105.0408
15.7194 | -105.0257
15.7232 | -105.0092
15.7269 | -104.9934
15.7307 | -104.9783
15.7344 | -104.9667
15.7382 | -104.9586
15.7419 | -104.9486
15.7457 | -104.9404
15.7494 | -104.934
15.7532 | -104.928
15.757 | -104.9229
15.7607 | -104.9196
15.7645 | -104.9198
15.7682 | -104.9225
15.772 | -104.9241
15.7757 | -104.9273
15.7795 | -104.932
15.7832 | -104.9387
15.787 | -104.9492
15.7907 | -104.9577
15.7945 | -104.9652
15.7982 | -104.9723
15.802 | -104.9789
15.8057 | -104.9864
15.8095 | -104.9918
15.8132 | -104.9987
15.817 | -105.0036
15.8207 | -105.0085
15.8245 | -105.0147
15.8282 | -105.0187
15.832 | -105.0215
15.8357 | -105.0234
15.8395 | -105.0222
15.8432 | -105.0201
15.847 | -105.0209
15.8507 | -105.0213
15.8545 | -105.0226
15.8582 | -105.0229
15.862 | -105.0222
15.8657 | -105.0249
15.8695 | -105.0286
15.8732 | -105.0315
15.877 | -105.0352
15.8807 | -105.0373
15.8845 | -105.0421
15.8882 | -105.054
15.892 | -105.0672
15.8957 | -105.0737
15.8995 | -105.0789
15.9032 | -105.0862
15.907 | -105.0968
15.9107 | -105.1081
15.9145 | -105.1211
15.9182 | -105.1346
15.922 | -105.1461
15.9257 | -105.1603
15.9295 | -105.1742
15.9332 | -105.1861
15.937 | -105.1979
15.9407 | -105.2096
15.9445 | -105.22
15.9482 | -105.2317
15.952 | -105.2434
15.9557 | -105.2535
15.9595 | -105.2607
15.9632 | -105.2665
15.967 | -105.2709
15.9707 | -105.2735
15.9745 | -105.273
15.9782 | -105.2731
15.982 | -105.2761
15.9857 | -105.2785
15.9895 | -105.2784
15.9932 | -105.2736
15.997 | -105.2633
16.0008 | -105.2537
16.0045 | -105.2492
16.0083 | -105.2425
16.012 | -105.2292
16.0158 | -105.2157
16.0195 | -105.2047
16.0233 | -105.1888
16.027 | -105.1742
16.0308 | -105.1599
16.0345 | -105.1418
16.0383 | -105.1237
16.042 | -105.1059
16.0458 | -105.0894
16.0495 | -105.0707
16.0533 | -105.053
16.057 | -105.0358
16.0608 | -105.0152
16.0645 | -104.9947
16.0683 | -104.9745
16.072 | -104.957
16.0758 | -104.9386
16.0795 | -104.9184
16.0833 | -104.9006
16.087 | -104.8841
16.0908 | -104.8689
16.0945 | -104.8537
16.0983 | -104.8401
16.102 | -104.8291
16.1058 | -104.8175
16.1095 | -104.8064
16.1133 | -104.7961
16.117 | -104.7868
16.1208 | -104.7811
16.1245 | -104.781
16.1283 | -104.7847
16.132 | -104.7892
16.1358 | -104.794
16.1395 | -104.8004
16.1433 | -104.811
16.147 | -104.8227
16.1508 | -104.8377
16.1545 | -104.854
16.1583 | -104.8722
16.162 | -104.8919
16.1658 | -104.9112
16.1695 | -104.9288
16.1733 | -104.9496
16.177 | -104.9677
16.1808 | -104.9813
16.1845 | -104.9975
16.1883 | -105.0143
16.192 | -105.0285
16.1958 | -105.0423
16.1995 | -105.0521
16.2033 | -105.0607
16.207 | -105.0639
16.2108 | -105.0643
16.2145 | -105.062
16.2183 | -105.0574
16.222 | -105.0539
16.2258 | -105.053
16.2295 | -105.0504
16.2333 | -105.0399
16.237 | -105.0308
16.2408 | -105.0183
16.2446 | -105.0089
16.2483 | -104.9991
16.2521 | -104.9863
16.2558 | -104.9783
16.2596 | -104.9677
16.2633 | -104.956
16.2671 | -104.9453
16.2708 | -104.933
16.2746 | -104.9206
16.2783 | -104.9082
16.2821 | -104.8974
16.2858 | -104.8831
16.2896 | -104.8696
16.2933 | -104.8606
16.2971 | -104.8518
16.3008 | -104.8435
16.3046 | -104.8334
16.3083 | -104.8216
16.3121 | -104.8108
16.3158 | -104.8001
16.3196 | -104.7904
16.3233 | -104.7815
16.3271 | -104.7724
16.3308 | -104.7616
16.3346 | -104.7509
16.3383 | -104.741
16.3421 | -104.7309
16.3458 | -104.7232
16.3496 | -104.7156
16.3533 | -104.7093
16.3571 | -104.7033
16.3608 | -104.696
16.3646 | -104.6895
16.3683 | -104.6842
16.3721 | -104.6791
16.3758 | -104.6757
16.3796 | -104.6767
16.3833 | -104.6769
16.3871 | -104.6768
16.3908 | -104.6792
16.3946 | -104.6827
16.3983 | -104.6892
16.4021 | -104.6993
16.4058 | -104.7071
16.4096 | -104.7174
16.4133 | -104.7326
16.4171 | -104.7446
16.4208 | -104.7583
16.4246 | -104.7751
16.4283 | -104.7902
16.4321 | -104.8059
16.4358 | -104.8207
16.4396 | -104.8411
16.4433 | -104.8595
16.4471 | -104.8794
16.4508 | -104.8978
16.4546 | -104.9139
16.4583 | -104.9294
16.4621 | -104.944
16.4658 | -104.9582
16.4696 | -104.9724
16.4733 | -104.9879
16.4771 | -105.0041
16.4808 | -105.0177
16.4846 | -105.0306
16.4883 | -105.0394
16.4921 | -105.0435
16.4959 | -105.0483
16.4996 | -105.0511
16.5034 | -105.0525
16.5071 | -105.0534
16.5109 | -105.0508
16.5146 | -105.0391
16.5184 | -105.0254
16.5221 | -105.0138
16.5259 | -105.0031
16.5296 | -104.9886
16.5334 | -104.972
16.5371 | -104.9516
16.5409 | -104.9293
16.5446 | -104.9101
16.5484 | -104.8882
16.5521 | -104.8628
16.5559 | -104.8391
16.5596 | -104.8155
16.5634 | -104.7915
16.5671 | -104.7679
16.5709 | -104.7448
16.5746 | -104.7194
16.5784 | -104.6961
16.5821 | -104.6728
16.5859 | -104.6514
16.5896 | -104.6318
16.5934 | -104.6131
16.5971 | -104.5982
16.6009 | -104.5854
16.6046 | -104.5725
16.6084 | -104.5622
16.6121 | -104.5544
16.6159 | -104.549
16.6196 | -104.5455
16.6234 | -104.5439
16.6271 | -104.545
16.6309 | -104.5508
16.6346 | -104.556
16.6384 | -104.5596
16.6421 | -104.5675
16.6459 | -104.5708
16.6496 | -104.5768
16.6534 | -104.5878
16.6571 | -104.5965
16.6609 | -104.6024
16.6646 | -104.6114
16.6684 | -104.6215
16.6721 | -104.6299
16.6759 | -104.6319
16.6796 | -104.6345
16.6834 | -104.6364
16.6871 | -104.6367
16.6909 | -104.6369
16.6946 | -104.6335
16.6984 | -104.6259
16.7021 | -104.6168
16.7059 | -104.6063
16.7096 | -104.5939
16.7134 | -104.5808
16.7171 | -104.5669
16.7209 | -104.5511
16.7246 | -104.5371
16.7284 | -104.5244
16.7321 | -104.5094
16.7359 | -104.4943
16.7397 | -104.4811
16.7434 | -104.4694
16.7472 | -104.4597
16.7509 | -104.4565
16.7547 | -104.4536
16.7584 | -104.4517
16.7622 | -104.4491
16.7659 | -104.4475
16.7697 | -104.4463
16.7734 | -104.4605
16.7772 | -104.4664
16.7809 | -104.4746
16.7847 | -104.4836
16.7884 | -104.494
16.7922 | -104.5074
16.7959 | -104.5188
16.7997 | -104.5308
16.8034 | -104.5458
16.8072 | -104.5621
16.8109 | -104.5773
16.8147 | -104.5911
16.8184 | -104.605
16.8222 | -104.6191
16.8259 | -104.6311
16.8297 | -104.6418
16.8334 | -104.6549
16.8372 | -104.667
16.8409 | -104.6805
16.8447 | -104.6947
16.8484 | -104.7075
16.8522 | -104.7193
16.8559 | -104.7303
16.8597 | -104.7429
16.8634 | -104.755
16.8672 | -104.7653
16.8709 | -104.7765
16.8747 | -104.7896
16.8784 | -104.804
16.8822 | -104.8167
16.8859 | -104.8337
16.8897 | -104.8471
16.8934 | -104.8545
16.8972 | -104.8656
16.9009 | -104.8787
16.9047 | -104.8895
16.9084 | -104.9032
16.9122 | -104.9157
16.9159 | -104.929
16.9197 | -104.9414
16.9234 | -104.953
16.9272 | -104.9662
16.9309 | -104.9781
16.9347 | -104.9889
16.9384 | -105.0009
16.9422 | -105.0137
16.9459 | -105.0246
16.9497 | -105.0315
16.9534 | -105.0352
16.9572 | -105.0341
16.9609 | -105.0294
16.9647 | -105.0223
16.9684 | -105.0146
16.9722 | -105.0083
16.9759 | -104.9994
16.9797 | -104.9851
16.9835 | -104.9657
16.9872 | -104.9459
16.991 | -104.9183
16.9947 | -104.8871
16.9985 | -104.8555
17.0022 | -104.8228
17.006 | -104.7866
17.0097 | -104.7494
17.0135 | -104.7043
17.0172 | -104.6546
17.021 | -104.609
17.0247 | -104.5608
17.0285 | -104.5073
17.0322 | -104.4506
17.036 | -104.3969
17.0397 | -104.3467
17.0435 | -104.297
17.0472 | -104.2472
17.051 | -104.1949
17.0547 | -104.1414
17.0585 | -104.0914
17.0622 | -104.0468
17.066 | -104.0028
17.0697 | -103.9591
17.0735 | -103.9124
17.0772 | -103.8667
17.081 | -103.8249
17.0847 | -103.7845
17.0885 | -103.7474
17.0922 | -103.7124
17.096 | -103.6813
17.0997 | -103.6528
17.1035 | -103.6251
17.1072 | -103.6009
17.111 | -103.5811
17.1147 | -103.5616
17.1185 | -103.5444
17.1222 | -103.5272
17.126 | -103.5107
17.1297 | -103.4952
17.1335 | -103.4827
17.1372 | -103.4736
17.141 | -103.4648
17.1447 | -103.4579
17.1485 | -103.4531
17.1522 | -103.4484
17.156 | -103.4448
17.1597 | -103.4416
17.1635 | -103.441
17.1672 | -103.4418
17.171 | -103.4432
17.1747 | -103.4462
17.1785 | -103.4527
17.1822 | -103.4612
17.186 | -103.4684
17.1897 | -103.4801
17.1935 | -103.4945
17.1972 | -103.5107
17.201 | -103.5321
17.2047 | -103.556
17.2085 | -103.5804
17.2122 | -103.6044
17.216 | -103.6262
17.2197 | -103.652
17.2235 | -103.6807
17.2272 | -103.7128
17.231 | -103.7431
17.2348 | -103.7756
17.2385 | -103.808
17.2423 | -103.8387
17.246 | -103.8694
17.2498 | -103.9018
17.2535 | -103.9334
17.2573 | -103.9689
17.261 | -103.9988
17.2648 | -104.0182
17.2685 | -104.0367
17.2723 | -104.0576
17.276 | -104.08
17.2798 | -104.0942
17.2835 | -104.1021
17.2873 | -104.1089
17.291 | -104.1154
17.2948 | -104.1229
17.2985 | -104.1286
17.3023 | -104.1286
17.306 | -104.1216
17.3098 | -104.1135
17.3135 | -104.0902
17.3173 | -104.0809
17.321 | -104.0743
17.3248 | -104.0637
17.3285 | -104.0484
17.3323 | -104.0311
17.336 | -104.0125
17.3398 | -103.9928
17.3435 | -103.9715
17.3473 | -103.9487
17.351 | -103.9277
17.3548 | -103.9058
17.3585 | -103.8817
17.3623 | -103.8597
17.366 | -103.8373
17.3698 | -103.8191
17.3735 | -103.8015
17.3773 | -103.7869
17.381 | -103.7712
17.3848 | -103.7538
17.3885 | -103.7387
17.3923 | -103.7239
17.396 | -103.7095
17.3998 | -103.6964
17.4035 | -103.6862
17.4073 | -103.6778
17.411 | -103.6686
17.4148 | -103.6597
17.4185 | -103.653
17.4223 | -103.6471
17.426 | -103.6394
17.4298 | -103.63
17.4335 | -103.62
17.4373 | -103.6085
17.441 | -103.6014
17.4448 | -103.5961
17.4485 | -103.5887
17.4523 | -103.5823
17.456 | -103.5758
17.4598 | -103.5698
17.4635 | -103.5645
17.4673 | -103.5574
17.471 | -103.551
17.4748 | -103.5448
17.4785 | -103.5399
17.4823 | -103.5378
17.4861 | -103.5342
17.4898 | -103.5279
17.4936 | -103.5233
17.4973 | -103.5208
17.5011 | -103.5192
17.5048 | -103.5172
17.5086 | -103.5177
17.5123 | -103.5191
17.5161 | -103.5211
17.5198 | -103.5252
17.5236 | -103.5309
17.5273 | -103.5386
17.5311 | -103.5465
17.5348 | -103.5564
17.5386 | -103.5674
17.5423 | -103.5805
17.5461 | -103.5958
17.5498 | -103.6144
17.5536 | -103.6348
17.5573 | -103.6586
17.5611 | -103.6832
17.5648 | -103.7111
17.5686 | -103.7377
17.5723 | -103.7668
17.5761 | -103.7977
17.5798 | -103.8282
17.5836 | -103.8578
17.5873 | -103.8873
17.5911 | -103.9173
17.5948 | -103.9457
17.5986 | -103.9747
17.6023 | -104.0051
17.6061 | -104.034
17.6098 | -104.0611
17.6136 | -104.0877
17.6173 | -104.1113
17.6211 | -104.1323
17.6248 | -104.1544
17.6286 | -104.1734
17.6323 | -104.1912
17.6361 | -104.2116
17.6398 | -104.2297
17.6436 | -104.2494
17.6473 | -104.269
17.6511 | -104.2877
17.6548 | -104.3043
17.6586 | -104.319
17.6623 | -104.3337
17.6661 | -104.3492
17.6698 | -104.3628
17.6736 | -104.3767
17.6773 | -104.392
17.6811 | -104.4028
17.6848 | -104.4125
17.6886 | -104.4209
17.6923 | -104.4256
17.6961 | -104.4288
17.6998 | -104.4327
17.7036 | -104.4339
17.7073 | -104.4342
17.7111 | -104.4324
17.7148 | -104.4294
17.7186 | -104.4262
17.7223 | -104.4209
17.7261 | -104.4152
17.7299 | -104.4056
17.7336 | -104.3944
17.7374 | -104.3854
17.7411 | -104.377
17.7449 | -104.3645
17.7486 | -104.3518
17.7524 | -104.3403
17.7561 | -104.3273
17.7599 | -104.3135
17.7636 | -104.3026
17.7674 | -104.2911
17.7711 | -104.2773
17.7749 | -104.262
17.7786 | -104.2487
17.7824 | -104.2368
17.7861 | -104.2238
17.7899 | -104.2128
17.7936 | -104.2042
17.7974 | -104.1986
17.8011 | -104.1927
17.8049 | -104.1878
17.8086 | -104.1836
17.8124 | -104.1821
17.8161 | -104.1842
17.8199 | -104.1906
17.8236 | -104.196
17.8274 | -104.2022
17.8311 | -104.2098
17.8349 | -104.2213
17.8386 | -104.2348
17.8424 | -104.2503
17.8461 | -104.2662
17.8499 | -104.2836
17.8536 | -104.303
17.8574 | -104.3175
17.8611 | -104.3346
17.8649 | -104.3537
17.8686 | -104.3712
17.8724 | -104.3931
17.8761 | -104.4139
17.8799 | -104.4309
17.8836 | -104.4475
17.8874 | -104.4652
17.8911 | -104.4805
17.8949 | -104.4961
17.8986 | -104.5094
17.9024 | -104.5232
17.9061 | -104.5375
17.9099 | -104.5506
17.9136 | -104.5644
17.9174 | -104.5777
17.9211 | -104.591
17.9249 | -104.6046
17.9286 | -104.6158
17.9324 | -104.6283
17.9361 | -104.6437
17.9399 | -104.6605
17.9436 | -104.6745
17.9474 | -104.6836
17.9511 | -104.6894
17.9549 | -104.6956
17.9586 | -104.7028
17.9624 | -104.7096
17.9661 | -104.7177
17.9699 | -104.7256
17.9737 | -104.7268
17.9774 | -104.7259
17.9812 | -104.7261
17.9849 | -104.7313
17.9887 | -104.7339
17.9924 | -104.7324
17.9962 | -104.7326
17.9999 | -104.7313
18.0037 | -104.7275
18.0074 | -104.7248
18.0112 | -104.7211
18.0149 | -104.7165
18.0187 | -104.7126
18.0224 | -104.7069
18.0262 | -104.6989
18.0299 | -104.692
18.0337 | -104.6858
18.0374 | -104.6816
18.0412 | -104.677
18.0449 | -104.6703
18.0487 | -104.6618
18.0524 | -104.6539
18.0562 | -104.6492
18.0599 | -104.6448
18.0637 | -104.6395
18.0674 | -104.6346
18.0712 | -104.6257
18.0749 | -104.6156
18.0787 | -104.6071
18.0824 | -104.6024
18.0862 | -104.6006
18.0899 | -104.598
18.0937 | -104.5937
18.0974 | -104.589
18.1012 | -104.5883
18.1049 | -104.5883
18.1087 | -104.5905
18.1124 | -104.5907
18.1162 | -104.5883
18.1199 | -104.5862
18.1237 | -104.5854
18.1274 | -104.5884
18.1312 | -104.5907
18.1349 | -104.5935
18.1387 | -104.5964
18.1424 | -104.5966
18.1462 | -104.5986
18.1499 | -104.6042
18.1537 | -104.6087
18.1574 | -104.6128
18.1612 | -104.6147
18.1649 | -104.614
18.1687 | -104.6144
18.1724 | -104.6148
18.1762 | -104.6154
18.1799 | -104.6151
18.1837 | -104.6144
18.1874 | -104.6146
18.1912 | -104.6135
18.1949 | -104.6122
18.1987 | -104.6101
18.2024 | -104.6072
18.2062 | -104.6073
18.2099 | -104.6091
18.2137 | -104.6087
18.2174 | -104.6073
18.2212 | -104.6068
18.225 | -104.6067
18.2287 | -104.6073
18.2325 | -104.61
18.2362 | -104.6119
18.24 | -104.6128
18.2437 | -104.6115
18.2475 | -104.6118
18.2512 | -104.6165
18.255 | -104.6192
18.2587 | -104.6217
18.2625 | -104.6248
18.2662 | -104.6267
18.27 | -104.6323
18.2737 | -104.6403
18.2775 | -104.6474
18.2812 | -104.6534
18.285 | -104.6587
18.2887 | -104.663
18.2925 | -104.6656
18.2962 | -104.6693
18.3 | -104.6719
18.3037 | -104.6724
18.3075 | -104.6712
18.3112 | -104.6676
18.315 | -104.6608
18.3187 | -104.6518
18.3225 | -104.6454
18.3262 | -104.6411
18.33 | -104.6341
18.3337 | -104.6229
18.3375 | -104.6075
18.3412 | -104.5886
18.345 | -104.5661
18.3487 | -104.5422
18.3525 | -104.522
18.3562 | -104.499
18.36 | -104.4725
18.3637 | -104.4463
18.3675 | -104.4198
18.3712 | -104.3972
18.375 | -104.3762
18.3787 | -104.3521
18.3825 | -104.3269
18.3862 | -104.2998
18.39 | -104.2733
18.3937 | -104.2495
18.3975 | -104.2314
18.4012 | -104.2115
18.405 | -104.1944
18.4087 | -104.1751
18.4125 | -104.1537
18.4162 | -104.1353
18.42 | -104.1204
18.4237 | -104.1076
18.4275 | -104.0995
18.4312 | -104.0929
18.435 | -104.0848
18.4387 | -104.0792
18.4425 | -104.0749
18.4462 | -104.0729
18.45 | -104.0734
18.4537 | -104.0737
18.4575 | -104.0733
18.4612 | -104.0763
18.465 | -104.0836
18.4688 | -104.0922
18.4725 | -104.1009
18.4763 | -104.1102
18.48 | -104.119
18.4838 | -104.1298
18.4875 | -104.1402
18.4913 | -104.1508
18.495 | -104.1625
18.4988 | -104.1722
18.5025 | -104.1748
18.5063 | -104.175
18.51 | -104.1791
18.5138 | -104.1832
18.5175 | -104.1845
18.5213 | -104.1847
18.525 | -104.1793
18.5288 | -104.17
18.5325 | -104.1594
18.5363 | -104.1497
18.54 | -104.1395
18.5438 | -104.1271
18.5475 | -104.1127
18.5513 | -104.0962
18.555 | -104.0799
18.5588 | -104.0602
18.5625 | -104.0419
18.5663 | -104.023
18.57 | -104.003
18.5738 | -103.9843
18.5775 | -103.9645
18.5813 | -103.9451
18.585 | -103.9239
18.5888 | -103.9018
18.5925 | -103.8832
18.5963 | -103.8649
18.6 | -103.8458
18.6038 | -103.8288
18.6075 | -103.8149
18.6113 | -103.8026
18.615 | -103.7881
18.6188 | -103.7763
18.6225 | -103.7629
18.6263 | -103.7528
18.63 | -103.7478
18.6338 | -103.7435
18.6375 | -103.7431
18.6413 | -103.7467
18.645 | -103.752
18.6488 | -103.7576
18.6525 | -103.7652
18.6563 | -103.7748
18.66 | -103.7894
18.6638 | -103.8071
18.6675 | -103.8249
18.6713 | -103.8432
18.675 | -103.8627
18.6788 | -103.8848
18.6825 | -103.907
18.6863 | -103.9282
18.69 | -103.9521
18.6938 | -103.9749
18.6975 | -103.9969
18.7013 | -104.0198
18.705 | -104.0419
18.7088 | -104.0648
18.7126 | -104.0864
18.7163 | -104.1067
18.7201 | -104.1285
18.7238 | -104.1483
18.7276 | -104.1661
18.7313 | -104.1834
18.7351 | -104.1992
18.7388 | -104.2149
18.7426 | -104.2304
18.7463 | -104.2452
18.7501 | -104.2601
18.7538 | -104.2774
18.7576 | -104.2934
18.7613 | -104.3084
18.7651 | -104.3259
18.7688 | -104.3476
18.7726 | -104.368
18.7763 | -104.3799
18.7801 | -104.3895
18.7838 | -104.4022
18.7876 | -104.4167
18.7913 | -104.432
18.7951 | -104.4466
18.7988 | -104.4565
18.8026 | -104.4628
18.8063 | -104.467
18.8101 | -104.4719
18.8138 | -104.4794
18.8176 | -104.4853
18.8213 | -104.4867
18.8251 | -104.4893
18.8288 | -104.4885
18.8326 | -104.4885
18.8363 | -104.4867
18.8401 | -104.4798
18.8438 | -104.4693
18.8476 | -104.4569
18.8513 | -104.4469
18.8551 | -104.4356
18.8588 | -104.4232
18.8626 | -104.4133
18.8663 | -104.4012
18.8701 | -104.3818
18.8738 | -104.3628
18.8776 | -104.3459
18.8813 | -104.3275
18.8851 | -104.3075
18.8888 | -104.2891
18.8926 | -104.2651
18.8963 | -104.2375
18.9001 | -104.2087
18.9038 | -104.1787
18.9076 | -104.1515
18.9113 | -104.1225
18.9151 | -104.0901
18.9188 | -104.0586
18.9226 | -104.0268
18.9263 | -103.992
18.9301 | -103.9547
18.9338 | -103.9188
18.9376 | -103.879
18.9413 | -103.8444
18.9451 | -103.8096
18.9488 | -103.7734
18.9526 | -103.7351
18.9564 | -103.6923
18.9601 | -103.6481
18.9639 | -103.606
18.9676 | -103.5678
18.9714 | -103.5324
18.9751 | -103.4953
18.9789 | -103.4588
18.9826 | -103.4238
18.9864 | -103.3896
18.9901 | -103.3561
18.9939 | -103.3286
18.9976 | -103.3002
19.0014 | -103.2711
19.0051 | -103.2396
19.0089 | -103.2085
19.0126 | -103.1795
19.0164 | -103.1533
19.0201 | -103.1288
19.0239 | -103.1017
19.0276 | -103.0745
19.0314 | -103.0491
19.0351 | -103.0247
19.0389 | -103.0028
19.0426 | -102.9838
19.0464 | -102.9661
19.0501 | -102.9482
19.0539 | -102.9298
19.0576 | -102.913
19.0614 | -102.8986
19.0651 | -102.8843
19.0689 | -102.8728
19.0726 | -102.8632
19.0764 | -102.8547
19.0801 | -102.8409
19.0839 | -102.8196
19.0876 | -102.7999
19.0914 | -102.7864
19.0951 | -102.7727
19.0989 | -102.7595
19.1026 | -102.7446
19.1064 | -102.7295
19.1101 | -102.7145
19.1139 | -102.7043
19.1176 | -102.6941
19.1214 | -102.6803
19.1251 | -102.6632
19.1289 | -102.6447
19.1326 | -102.6261
19.1364 | -102.6102
19.1401 | -102.5944
19.1439 | -102.5761
19.1476 | -102.5538
19.1514 | -102.5284
19.1551 | -102.5022
19.1589 | -102.4773
19.1626 | -102.4544
19.1664 | -102.432
19.1701 | -102.4086
19.1739 | -102.3845
19.1776 | -102.3584
19.1814 | -102.3346
19.1851 | -102.3133
19.1889 | -102.2918
19.1926 | -102.2714
19.1964 | -102.2517
19.2002 | -102.2336
19.2039 | -102.2151
19.2077 | -102.1981
19.2114 | -102.1831
19.2152 | -102.1682
19.2189 | -102.1526
19.2227 | -102.1386
19.2264 | -102.1286
19.2302 | -102.1233
19.2339 | -102.1208
19.2377 | -102.1175
19.2414 | -102.1131
19.2452 | -102.1095
19.2489 | -102.1048
19.2527 | -102.1003
19.2564 | -102.0969
19.2602 | -102.0954
19.2639 | -102.0954
19.2677 | -102.0941
19.2714 | -102.0924
19.2752 | -102.09
19.2789 | -102.0917
19.2827 | -102.0956
19.2864 | -102.0979
19.2902 | -102.1
19.2939 | -102.1001
19.2977 | -102.0997
19.3014 | -102.0973
19.3052 | -102.0967
19.3089 | -102.0933
19.3127 | -102.0909
19.3164 | -102.0863
19.3202 | -102.0765
19.3239 | -102.0615
19.3277 | -102.0431
19.3314 | -102.0216
19.3352 | -102.0057
19.3389 | -101.9936
19.3427 | -101.9763
19.3464 | -101.9504
19.3502 | -101.922
19.3539 | -101.8979
19.3577 | -101.8787
19.3614 | -101.859
19.3652 | -101.8319
19.3689 | -101.8032
19.3727 | -101.7737
19.3764 | -101.7433
19.3802 | -101.7125
19.3839 | -101.6811
19.3877 | -101.6476
19.3914 | -101.6119
19.3952 | -101.5753
19.3989 | -101.5405
19.4027 | -101.5052
19.4064 | -101.4695
19.4102 | -101.4327
19.4139 | -101.3976
19.4177 | -101.3638
19.4214 | -101.3313
19.4252 | -101.302
19.4289 | -101.2753
19.4327 | -101.2495
19.4364 | -101.2239
19.4402 | -101.1997
19.444 | -101.1795
19.4477 | -101.1604
19.4515 | -101.1455
19.4552 | -101.1351
19.459 | -101.1282
19.4627 | -101.1231
19.4665 | -101.1212
19.4702 | -101.1243
19.474 | -101.133
19.4777 | -101.1704
19.4815 | -101.1878
19.4852 | -101.2091
19.489 | -101.2355
19.4927 | -101.2654
19.4965 | -101.3005
19.5002 | -101.3388
19.504 | -101.3782
19.5077 | -101.4202
19.5115 | -101.4636
19.5152 | -101.5094
19.519 | -101.558
19.5227 | -101.6085
19.5265 | -101.6644
19.5302 | -101.7214
19.534 | -101.7755
19.5377 | -101.8289
19.5415 | -101.882
19.5452 | -101.9361
19.549 | -101.9905
19.5527 | -102.0471
19.5565 | -102.0998
19.5602 | -102.1435
19.564 | -102.1783
19.5677 | -102.2085
19.5715 | -102.2412
19.5752 | -102.2779
19.579 | -102.3144
19.5827 | -102.3379
19.5865 | -102.3524
19.5902 | -102.3563
19.594 | -102.3607
19.5977 | -102.3703
19.6015 | -102.3808
19.6052 | -102.3865
19.609 | -102.3861
19.6127 | -102.3716
19.6165 | -102.3475
19.6202 | -102.3229
19.624 | -102.3017
19.6277 | -102.2772
19.6315 | -102.2425
19.6352 | -102.1964
19.639 | -102.141
19.6427 | -102.085
19.6465 | -102.0318
19.6502 | -101.9791
19.654 | -101.9197
19.6577 | -101.8495
19.6615 | -101.7715
19.6652 | -101.6914
19.669 | -101.6214
19.6727 | -101.5582
19.6765 | -101.4952
19.6802 | -101.4299
19.684 | -101.3637
19.6877 | -101.3001
19.6915 | -101.2389
19.6953 | -101.1802
19.699 | -101.1243
19.7028 | -101.0698
19.7065 | -101.0153
19.7103 | -100.9608
19.714 | -100.9088
19.7178 | -100.8609
19.7215 | -100.8175
19.7253 | -100.7782
19.729 | -100.7382
19.7328 | -100.6969
19.7365 | -100.6602
19.7403 | -100.6298
19.744 | -100.6029
19.7478 | -100.5784
19.7515 | -100.5568
19.7553 | -100.538
19.759 | -100.5217
19.7628 | -100.5084
19.7665 | -100.4978
19.7703 | -100.4903
19.774 | -100.488
19.7778 | -100.4883
19.7815 | -100.4896
19.7853 | -100.495
19.789 | -100.5049
19.7928 | -100.5188
19.7965 | -100.5376
19.8003 | -100.5604
19.804 | -100.5869
19.8078 | -100.6156
19.8115 | -100.6477
19.8153 | -100.6858
19.819 | -100.7282
19.8228 | -100.7738
19.8265 | -100.8206
19.8303 | -100.8708
19.834 | -100.923
19.8378 | -100.9796
19.8415 | -101.0426
19.8453 | -101.1119
19.849 | -101.1827
19.8528 | -101.2507
19.8565 | -101.3163
19.8603 | -101.3821
19.864 | -101.4547
19.8678 | -101.5309
19.8715 | -101.6085
19.8753 | -101.6817
19.879 | -101.7498
19.8828 | -101.8159
19.8865 | -101.8823
19.8903 | -101.9477
19.894 | -102.0147
19.8978 | -102.0825
19.9015 | -102.1473
19.9053 | -102.2079
19.909 | -102.2669
19.9128 | -102.3277
19.9165 | -102.3906
19.9203 | -102.4535
19.924 | -102.5111
19.9278 | -102.555
19.9315 | -102.5923
19.9353 | -102.6353
19.9391 | -102.6857
19.9428 | -102.7378
19.9466 | -102.7808
19.9503 | -102.8013
19.9541 | -102.8083
19.9578 | -102.811
19.9616 | -102.8195
19.9653 | -102.8274
19.9691 | -102.827
19.9728 | -102.8187
19.9766 | -102.8042
19.9803 | -102.7857
19.9841 | -102.7664
19.9878 | -102.7457
19.9916 | -102.724
19.9953 | -102.6967
19.9991 | -102.6608
20.0028 | -102.6204
20.0066 | -102.5786
20.0103 | -102.537
20.0141 | -102.4945
20.0178 | -102.4225
20.0216 | -102.3731
20.0253 | -102.3197
20.0291 | -102.263
20.0328 | -102.2054
20.0366 | -102.1479
20.0403 | -102.0921
20.0441 | -102.0357
20.0478 | -101.9788
20.0516 | -101.9227
20.0553 | -101.8689
20.0591 | -101.816
20.0628 | -101.7655
20.0666 | -101.7185
20.0703 | -101.6727
20.0741 | -101.6277
20.0778 | -101.5851
20.0816 | -101.546
20.0853 | -101.5096
20.0891 | -101.4774
20.0928 | -101.4497
20.0966 | -101.4256
20.1003 | -101.4057
20.1041 | -101.3906
20.1078 | -101.3789
20.1116 | -101.3722
20.1153 | -101.3702
20.1191 | -101.3714
20.1228 | -101.3766
20.1266 | -101.386
20.1303 | -101.4005
20.1341 | -101.4186
20.1378 | -101.4401
20.1416 | -101.4654
20.1453 | -101.4952
20.1491 | -101.5282
20.1528 | -101.5639
20.1566 | -101.6027
20.1603 | -101.6436
20.1641 | -101.688
20.1678 | -101.7349
20.1716 | -101.782
20.1753 | -101.8294
20.1791 | -101.8783
20.1828 | -101.9278
20.1866 | -101.9785
20.1904 | -102.03
20.1941 | -102.0809
20.1979 | -102.1318
20.2016 | -102.1806
20.2054 | -102.2264
20.2091 | -102.2727
20.2129 | -102.3208
20.2166 | -102.3668
20.2204 | -102.408
20.2241 | -102.4437
20.2279 | -102.4743
20.2316 | -102.5034
20.2354 | -102.5321
20.2391 | -102.5565
20.2429 | -102.5749
20.2466 | -102.5831
20.2504 | -102.5869
20.2541 | -102.5896
20.2579 | -102.5888
20.2616 | -102.5809
20.2654 | -102.5645
20.2691 | -102.5427
20.2729 | -102.5168
20.2766 | -102.4871
20.2804 | -102.4571
20.2841 | -102.4263
20.2879 | -102.39
20.2916 | -102.3512
20.2954 | -102.3122
20.2991 | -102.2717
20.3029 | -102.2312
20.3066 | -102.1912
20.3104 | -102.1518
20.3141 | -102.1105
20.3179 | -102.0697
20.3216 | -102.0308
20.3254 | -101.9934
20.3291 | -101.9576
20.3329 | -101.9258
20.3366 | -101.8972
20.3404 | -101.8713
20.3441 | -101.8483
20.3479 | -101.8283
20.3516 | -101.8128
20.3554 | -101.8021
20.3591 | -101.7956
20.3629 | -101.7935
20.3666 | -101.7955
20.3704 | -101.7997
20.3741 | -101.8083
20.3779 | -101.8209
20.3816 | -101.8383
20.3854 | -101.8615
20.3891 | -101.8875
20.3929 | -101.9164
20.3966 | -101.9501
20.4004 | -101.99
20.4041 | -102.033
20.4079 | -102.0789
20.4116 | -102.1292
20.4154 | -102.1831
20.4191 | -102.2397
20.4229 | -102.3004
20.4266 | -102.3643
20.4304 | -102.4315
20.4342 | -102.5009
20.4379 | -102.5712
20.4417 | -102.6445
20.4454 | -102.7175
20.4492 | -102.7911
20.4529 | -102.8627
20.4567 | -102.935
20.4604 | -103.0095
20.4642 | -103.0821
20.4679 | -103.1494
20.4717 | -103.2142
20.4754 | -103.2768
20.4792 | -103.3371
20.4829 | -103.3947
20.4867 | -103.4448
20.4904 | -103.4882
20.4942 | -103.527
20.4979 | -103.559
20.5017 | -103.5862
20.5054 | -103.6077
20.5092 | -103.6232
20.5129 | -103.6354
20.5167 | -103.6424
20.5204 | -103.6432
20.5242 | -103.6381
20.5279 | -103.6285
20.5317 | -103.6165
20.5354 | -103.6004
20.5392 | -103.5782
20.5429 | -103.5498
20.5467 | -103.5179
20.5504 | -103.4882
20.5542 | -103.4621
20.5579 | -103.4333
20.5617 | -103.394
20.5654 | -103.364
20.5692 | -103.3322
20.5729 | -103.2973
20.5767 | -103.263
20.5804 | -103.2294
20.5842 | -103.1962
20.5879 | -103.1662
20.5917 | -103.1402
20.5954 | -103.1156
20.5992 | -103.0915
20.6029 | -103.0675
20.6067 | -103.0474
20.6104 | -103.0323
20.6142 | -103.0217
20.6179 | -103.013
20.6217 | -103.0064
20.6254 | -103.0005
20.6292 | -102.9984
20.6329 | -103.0003
20.6367 | -103.0023
20.6404 | -103.0061
20.6442 | -103.0124
20.6479 | -103.0223
20.6517 | -103.0322
20.6554 | -103.0421
20.6592 | -103.0521
20.6629 | -103.0655
20.6667 | -103.0808
20.6704 | -103.0949
20.6742 | -103.1078
20.6779 | -103.1193
20.6817 | -103.1314
20.6855 | -103.1454
20.6892 | -103.1586
20.693 | -103.1693
20.6967 | -103.1779
20.7005 | -103.1864
20.7042 | -103.1938
20.708 | -103.2007
20.7117 | -103.206
20.7155 | -103.2091
20.7192 | -103.2095
20.723 | -103.2077
20.7267 | -103.2046
20.7305 | -103.2002
20.7342 | -103.1954
20.738 | -103.1893
20.7417 | -103.1811
20.7455 | -103.1738
20.7492 | -103.1631
20.753 | -103.1495
20.7567 | -103.1377
20.7605 | -103.1244
20.7642 | -103.1092
20.768 | -103.0929
20.7717 | -103.0744
20.7755 | -103.0538
20.7792 | -103.0362
20.783 | -103.0212
20.7867 | -103.0074
20.7905 | -102.9921
20.7942 | -102.9725
20.798 | -102.9543
20.8017 | -102.9389
20.8055 | -102.9241
20.8092 | -102.9127
20.813 | -102.9036
20.8167 | -102.8959
20.8205 | -102.8904
20.8242 | -102.8856
20.828 | -102.8817
20.8317 | -102.879
20.8355 | -102.8796
20.8392 | -102.8827
20.843 | -102.8862
20.8467 | -102.8902
20.8505 | -102.8964
20.8542 | -102.9054
20.858 | -102.9171
20.8617 | -102.9306
20.8655 | -102.9456
20.8692 | -102.9627
20.873 | -102.9807
20.8767 | -103.0019
20.8805 | -103.0252
20.8842 | -103.0466
20.888 | -103.0696
20.8917 | -103.0954
20.8955 | -103.1231
20.8992 | -103.153
20.903 | -103.1844
20.9067 | -103.216
20.9105 | -103.246
20.9142 | -103.2773
20.918 | -103.3116
20.9217 | -103.3481
20.9255 | -103.3828
20.9293 | -103.4175
20.933 | -103.4555
20.9368 | -103.4942
20.9405 | -103.5336
20.9443 | -103.5716
20.948 | -103.6129
20.9518 | -103.6555
20.9555 | -103.697
20.9593 | -103.7377
20.963 | -103.7768
20.9668 | -103.8156
20.9705 | -103.8547
20.9743 | -103.8945
20.978 | -103.9324
20.9818 | -103.9662
20.9855 | -103.9991
20.9893 | -104.0258
20.993 | -104.0519
20.9968 | -104.0803
21.0005 | -104.1057
21.0043 | -104.1272
21.008 | -104.1453
21.0118 | -104.1579
21.0155 | -104.1677
21.0193 | -104.1749
21.023 | -104.1815
21.0268 | -104.1864
21.0305 | -104.1882
21.0343 | -104.1852
21.038 | -104.1801
21.0418 | -104.1702
21.0455 | -104.1615
21.0493 | -104.1544
21.053 | -104.1423
21.0568 | -104.1297
21.0605 | -104.1152
21.0643 | -104.1014
21.068 | -104.0875
21.0718 | -104.0737
21.0755 | -104.0553
21.0793 | -104.0378
21.083 | -104.0187
21.0868 | -104.0001
21.0905 | -103.9869
21.0943 | -103.9728
21.098 | -103.9553
21.1018 | -103.9448
21.1055 | -103.9313
21.1093 | -103.9163
21.113 | -103.9014
21.1168 | -103.8877
21.1205 | -103.8762
21.1243 | -103.867
21.128 | -103.8602
21.1318 | -103.8546
21.1355 | -103.85
21.1393 | -103.8488
21.143 | -103.8496
21.1468 | -103.8527
21.1505 | -103.8561
21.1543 | -103.8617
21.158 | -103.8693
21.1618 | -103.8779
21.1655 | -103.8877
21.1693 | -103.897
21.1731 | -103.9085
21.1768 | -103.9227
21.1806 | -103.9371
21.1843 | -103.9533
21.1881 | -103.9725
21.1918 | -103.9909
21.1956 | -104.0111
21.1993 | -104.0314
21.2031 | -104.0536
21.2068 | -104.0744
21.2106 | -104.0959
21.2143 | -104.1189
21.2181 | -104.1439
21.2218 | -104.1678
21.2256 | -104.1862
21.2293 | -104.2028
21.2331 | -104.2202
21.2368 | -104.2399
21.2406 | -104.2591
21.2443 | -104.2787
21.2481 | -104.2929
21.2518 | -104.3034
21.2556 | -104.3122
21.2593 | -104.3186
21.2631 | -104.3275
21.2668 | -104.3382
21.2706 | -104.3446
21.2743 | -104.3448
21.2781 | -104.3399
21.2818 | -104.3325
21.2856 | -104.3251
21.2893 | -104.3216
21.2931 | -104.315
21.2968 | -104.3019
21.3006 | -104.2864
21.3043 | -104.2682
21.3081 | -104.25
21.3118 | -104.2339
21.3156 | -104.2165
21.3193 | -104.1982
21.3231 | -104.1747
21.3268 | -104.1444
21.3306 | -104.1144
21.3343 | -104.0881
21.3381 | -104.0647
21.3418 | -104.041
21.3456 | -104.0146
21.3493 | -103.9872
21.3531 | -103.9605
21.3568 | -103.9331
21.3606 | -103.9086
21.3643 | -103.8829
21.3681 | -103.8572
21.3718 | -103.832
21.3756 | -103.8114
21.3793 | -103.7885
21.3831 | -103.7666
21.3868 | -103.7464
21.3906 | -103.7268
21.3943 | -103.7072
21.3981 | -103.6882
21.4018 | -103.6703
21.4056 | -103.6531
21.4093 | -103.6364
21.4131 | -103.6226
21.4169 | -103.6115
21.4206 | -103.602
21.4244 | -103.5931
21.4281 | -103.586
21.4319 | -103.5809
21.4356 | -103.5772
21.4394 | -103.5753
21.4431 | -103.5745
21.4469 | -103.5752
21.4506 | -103.5775
21.4544 | -103.5825
21.4581 | -103.5887
21.4619 | -103.5968
21.4656 | -103.6057
21.4694 | -103.6172
21.4731 | -103.6322
21.4769 | -103.6458
21.4806 | -103.6591
21.4844 | -103.6766
21.4881 | -103.6936
21.4919 | -103.7074
21.4956 | -103.7233
21.4994 | -103.7369
21.5031 | -103.7511
21.5069 | -103.767
21.5106 | -103.7851
21.5144 | -103.8038
21.5181 | -103.8201
21.5219 | -103.8334
21.5256 | -103.8479
21.5294 | -103.8686
21.5331 | -103.8892
21.5369 | -103.9044
21.5406 | -103.916
21.5444 | -103.9275
21.5481 | -103.9385
21.5519 | -103.9506
21.5556 | -103.9651
21.5594 | -103.9806
21.5631 | -103.9925
21.5669 | -103.9995
21.5706 | -104.0053
21.5744 | -104.0117
21.5781 | -104.0214
21.5819 | -104.033
21.5856 | -104.0432
21.5894 | -104.05
21.5931 | -104.0535
21.5969 | -104.0562
21.6006 | -104.0594
21.6044 | -104.0643
21.6081 | -104.0698
21.6119 | -104.0741
21.6156 | -104.0763
21.6194 | -104.0778
21.6231 | -104.0792
21.6269 | -104.0843
21.6306 | -104.0933
21.6344 | -104.0994
21.6381 | -104.1013
21.6419 | -104.103
21.6456 | -104.1015
21.6494 | -104.1015
21.6531 | -104.1023
21.6569 | -104.1035
21.6607 | -104.1018
21.6644 | -104.0974
21.6682 | -104.0935
21.6719 | -104.0918
21.6757 | -104.0936
21.6794 | -104.0956
21.6832 | -104.0937
21.6869 | -104.0898
21.6907 | -104.0869
21.6944 | -104.0851
21.6982 | -104.0823
21.7019 | -104.081
21.7057 | -104.0809
21.7094 | -104.0771
21.7132 | -104.0714
21.7169 | -104.0642
21.7207 | -104.055
21.7244 | -104.0447
21.7282 | -104.0317
21.7319 | -104.0159
21.7357 | -104.0001
21.7394 | -103.9859
21.7432 | -103.9706
21.7469 | -103.953
21.7507 | -103.9335
21.7544 | -103.9153
21.7582 | -103.8958
21.7619 | -103.8752
21.7657 | -103.855
21.7694 | -103.8333
21.7732 | -103.8095
21.7769 | -103.7866
21.7807 | -103.763
21.7844 | -103.7399
21.7882 | -103.7165
21.7919 | -103.6919
21.7957 | -103.6712
21.7994 | -103.6545
21.8032 | -103.6383
21.8069 | -103.6239
21.8107 | -103.6082
21.8144 | -103.5901
21.8182 | -103.57
21.8219 | -103.5539
21.8257 | -103.5409
21.8294 | -103.529
21.8332 | -103.5193
21.8369 | -103.5084
21.8407 | -103.4989
21.8444 | -103.4907
21.8482 | -103.4847
21.8519 | -103.4802
21.8557 | -103.4766
21.8594 | -103.472
21.8632 | -103.4688
21.8669 | -103.4671
21.8707 | -103.4682
21.8744 | -103.4705
21.8782 | -103.4727
21.8819 | -103.476
21.8857 | -103.4778
21.8894 | -103.4784
21.8932 | -103.4798
21.8969 | -103.4818
21.9007 | -103.4874
21.9044 | -103.4928
21.9082 | -103.4986
21.912 | -103.4997
21.9157 | -103.5018
21.9195 | -103.5046
21.9232 | -103.5071
21.927 | -103.5094
21.9307 | -103.5124
21.9345 | -103.5144
21.9382 | -103.5163
21.942 | -103.5187
21.9457 | -103.5223
21.9495 | -103.5246
21.9532 | -103.5282
21.957 | -103.5322
21.9607 | -103.5325
21.9645 | -103.532
21.9682 | -103.5314
21.972 | -103.5345
21.9757 | -103.5379
21.9795 | -103.5381
21.9832 | -103.5349
21.987 | -103.5311
21.9907 | -103.5285
21.9945 | -103.5271
21.9982 | -103.5273
22.002 | -103.5257
22.0057 | -103.522
22.0095 | -103.5189
22.0132 | -103.5155
22.017 | -103.5077
22.0207 | -103.4995
22.0245 | -103.4943
22.0282 | -103.4909
22.032 | -103.489
22.0357 | -103.4861
22.0395 | -103.4841
22.0432 | -103.4812
22.047 | -103.4758
22.0507 | -103.468
22.0545 | -103.46
22.0582 | -103.4525
22.062 | -103.4476
22.0657 | -103.4401
22.0695 | -103.4324
22.0732 | -103.4226
22.077 | -103.4148
22.0807 | -103.4065
22.0845 | -103.4019
22.0882 | -103.3972
22.092 | -103.3929
22.0957 | -103.3874
22.0995 | -103.38
22.1032 | -103.3715
22.107 | -103.3628
22.1107 | -103.3542
22.1145 | -103.3446
22.1182 | -103.3357
22.122 | -103.3296
22.1257 | -103.3232
22.1295 | -103.3189
22.1332 | -103.3106
22.137 | -103.3021
22.1407 | -103.2944
22.1445 | -103.2839
22.1482 | -103.2744
22.152 | -103.2653
22.1558 | -103.257
22.1595 | -103.2486
22.1633 | -103.2409
22.167 | -103.2329
22.1708 | -103.2274
22.1745 | -103.2217
22.1783 | -103.2163
22.182 | -103.2105
22.1858 | -103.2092
22.1895 | -103.2061
22.1933 | -103.2042
22.197 | -103.2035
22.2008 | -103.2041
22.2045 | -103.2077
22.2083 | -103.2111
22.212 | -103.2141
22.2158 | -103.2139
22.2195 | -103.2126
22.2233 | -103.2112
22.227 | -103.2167
22.2308 | -103.2238
22.2345 | -103.2302
22.2383 | -103.235
22.242 | -103.2389
22.2458 | -103.2421
22.2495 | -103.2444
22.2533 | -103.2481
22.257 | -103.2538
22.2608 | -103.2594
22.2645 | -103.2658
22.2683 | -103.2721
22.272 | -103.2788
22.2758 | -103.2888
22.2795 | -103.2986
22.2833 | -103.3047
22.287 | -103.3079
22.2908 | -103.3129
22.2945 | -103.3194
22.2983 | -103.3262
22.302 | -103.3274
22.3058 | -103.3282
22.3095 | -103.3289
22.3133 | -103.3303
22.317 | -103.3311
22.3208 | -103.3315
22.3245 | -103.3309
22.3283 | -103.3297
22.332 | -103.3321
22.3358 | -103.3353
22.3395 | -103.3346
22.3433 | -103.3313
22.347 | -103.3256
22.3508 | -103.3213
22.3545 | -103.317
22.3583 | -103.3134
22.362 | -103.3091
22.3658 | -103.305
22.3695 | -103.3019
22.3733 | -103.2992
22.377 | -103.2943
22.3808 | -103.2902
22.3845 | -103.2867
22.3883 | -103.2846
22.392 | -103.2793
22.3958 | -103.2751
22.3996 | -103.2713
22.4033 | -103.2669
22.4071 | -103.2624
22.4108 | -103.2567
22.4146 | -103.252
22.4183 | -103.2482
22.4221 | -103.2445
22.4258 | -103.242
22.4296 | -103.2403
22.4333 | -103.2393
22.4371 | -103.2353
22.4408 | -103.2315
22.4446 | -103.2278
22.4483 | -103.2268
22.4521 | -103.2254
22.4558 | -103.2232
22.4596 | -103.2257
22.4633 | -103.2278
22.4671 | -103.2238
22.4708 | -103.2179
22.4746 | -103.2138
22.4783 | -103.2077
22.4821 | -103.2023
22.4858 | -103.1953
22.4896 | -103.1848
22.4933 | -103.1739
22.4971 | -103.1614
22.5008 | -103.1456
22.5046 | -103.1251
22.5083 | -103.107
22.5121 | -103.09
22.5158 | -103.0714
22.5196 | -103.0533
22.5233 | -103.0331
22.5271 | -103.0121
22.5308 | -102.9902
22.5346 | -102.9685
22.5383 | -102.9472
22.5421 | -102.9239
22.5458 | -102.8952
22.5496 | -102.8622
22.5533 | -102.8285
22.5571 | -102.7935
22.5608 | -102.7589
22.5646 | -102.7258
22.5683 | -102.6936
22.5721 | -102.6607
22.5758 | -102.626
22.5796 | -102.5922
22.5833 | -102.5584
22.5871 | -102.5232
22.5908 | -102.4888
22.5946 | -102.4565
22.5983 | -102.4256
22.6021 | -102.3986
22.6058 | -102.3739
22.6096 | -102.3494
22.6133 | -102.3251
22.6171 | -102.3029
22.6208 | -102.2851
22.6246 | -102.2681
22.6283 | -102.252
22.6321 | -102.2372
22.6358 | -102.226
22.6396 | -102.2178
22.6433 | -102.2123
22.6471 | -102.2088
22.6509 | -102.2073
22.6546 | -102.207
22.6584 | -102.2094
22.6621 | -102.2147
22.6659 | -102.221
22.6696 | -102.2306
22.6734 | -102.2429
22.6771 | -102.2553
22.6809 | -102.2703
22.6846 | -102.2886
22.6884 | -102.3073
22.6921 | -102.3265
22.6959 | -102.3459
22.6996 | -102.3656
22.7034 | -102.3866
22.7071 | -102.4074
22.7109 | -102.4285
22.7146 | -102.4526
22.7184 | -102.4774
22.7221 | -102.5011
22.7259 | -102.5212
22.7296 | -102.5432
22.7334 | -102.5667
22.7371 | -102.59
22.7409 | -102.6128
22.7446 | -102.6333
22.7484 | -102.6509
22.7521 | -102.6692
22.7559 | -102.6858
22.7596 | -102.7022
22.7634 | -102.7183
22.7671 | -102.7329
22.7709 | -102.7448
22.7746 | -102.756
22.7784 | -102.7663
22.7821 | -102.775
22.7859 | -102.7825
22.7896 | -102.7925
22.7934 | -102.8022
22.7971 | -102.8076
22.8009 | -102.81
22.8046 | -102.8113
22.8084 | -102.814
22.8121 | -102.8174
22.8159 | -102.817
22.8196 | -102.8165
22.8234 | -102.8195
22.8271 | -102.8222
22.8309 | -102.8178
22.8346 | -102.8127
22.8384 | -102.807
22.8421 | -102.7985
22.8459 | -102.7894
22.8496 | -102.7808
22.8534 | -102.7756
22.8571 | -102.7726
22.8609 | -102.7697
22.8646 | -102.7651
22.8684 | -102.7601
22.8721 | -102.7586
22.8759 | -102.7589
22.8796 | -102.7625
22.8834 | -102.7646
22.8871 | -102.7632
22.8909 | -102.7618
22.8946 | -102.7614
22.8984 | -102.7614
22.9022 | -102.7612
22.9059 | -102.762
22.9097 | -102.7633
22.9134 | -102.7636
22.9172 | -102.7646
22.9209 | -102.7668
22.9247 | -102.7687
22.9284 | -102.7736
22.9322 | -102.7785
22.9359 | -102.7833
22.9397 | -102.7883
22.9434 | -102.793
22.9472 | -102.7974
22.9509 | -102.8039
22.9547 | -102.8102
22.9584 | -102.8139
22.9622 | -102.8168
22.9659 | -102.8212
22.9697 | -102.8255
22.9734 | -102.8303
22.9772 | -102.8355
22.9809 | -102.8383
22.9847 | -102.8405
22.9884 | -102.8442
22.9922 | -102.8451
22.9959 | -102.8448
22.9997 | -102.847
23.0034 | -102.8493
23.0072 | -102.8482
23.0109 | -102.8477
23.0147 | -102.8472
23.0184 | -102.8449
23.0222 | -102.8423
23.0259 | -102.8413
23.0297 | -102.8424
23.0334 | -102.8396
23.0372 | -102.8302
23.0409 | -102.8162
23.0447 | -102.8037
23.0484 | -102.7917
23.0522 | -102.7803
23.0559 | -102.7697
23.0597 | -102.7603
23.0634 | -102.7469
23.0672 | -102.7295
23.0709 | -102.7127
23.0747 | -102.6955
23.0784 | -102.6778
23.0822 | -102.6601
23.0859 | -102.6386
23.0897 | -102.6124
23.0934 | -102.5852
23.0972 | -102.5582
23.1009 | -102.5285
23.1047 | -102.4982
23.1084 | -102.4652
23.1122 | -102.4333
23.1159 | -102.4017
23.1197 | -102.3679
23.1234 | -102.3343
23.1272 | -102.3013
23.1309 | -102.268
23.1347 | -102.2321
23.1384 | -102.1945
23.1422 | -102.1589
23.146 | -102.1243
23.1497 | -102.088
23.1535 | -102.0536
23.1572 | -102.0204
23.161 | -101.9844
23.1647 | -101.9495
23.1685 | -101.9172
23.1722 | -101.8865
23.176 | -101.8561
23.1797 | -101.8251
23.1835 | -101.7949
23.1872 | -101.7664
23.191 | -101.743
23.1947 | -101.7211
23.1985 | -101.6987
23.2022 | -101.6792
23.206 | -101.6611
23.2097 | -101.6451
23.2135 | -101.6296
23.2172 | -101.6137
23.221 | -101.5985
23.2247 | -101.5856
23.2285 | -101.5757
23.2322 | -101.5677
23.236 | -101.5625
23.2397 | -101.557
23.2435 | -101.5501
23.2472 | -101.5471
23.251 | -101.5447
23.2547 | -101.541
23.2585 | -101.5383
23.2622 | -101.5386
23.266 | -101.5439
23.2697 | -101.5434
23.2735 | -101.5434
23.2772 | -101.5442
23.281 | -101.5443
23.2847 | -101.5423
23.2885 | -101.5427
23.2922 | -101.5419
23.296 | -101.5394
23.2997 | -101.5379
23.3035 | -101.5337
23.3072 | -101.5282
23.311 | -101.5247
23.3147 | -101.5211
23.3185 | -101.5171
23.3222 | -101.5128
23.326 | -101.5086
23.3297 | -101.5036
23.3335 | -101.4976
23.3372 | -101.4897
23.341 | -101.4828
23.3447 | -101.4771
23.3485 | -101.4698
23.3522 | -101.4597
23.356 | -101.4471
23.3597 | -101.4346
23.3635 | -101.4266
23.3672 | -101.4206
23.371 | -101.4154
23.3747 | -101.4074
23.3785 | -101.3959
23.3822 | -101.386
23.386 | -101.3756
23.3898 | -101.3631
23.3935 | -101.351
23.3973 | -101.3404
23.401 | -101.3295
23.4048 | -101.3173
23.4085 | -101.3053
23.4123 | -101.2939
23.416 | -101.2842
23.4198 | -101.2776
23.4235 | -101.2697
23.4273 | -101.2604
23.431 | -101.2524
23.4348 | -101.2453
23.4385 | -101.2381
23.4423 | -101.2278
23.446 | -101.2152
23.4498 | -101.2043
23.4535 | -101.1957
23.4573 | -101.1883
23.461 | -101.1808
23.4648 | -101.1746
23.4685 | -101.1682
23.4723 | -101.1603
23.476 | -101.1525
23.4798 | -101.1463
23.4835 | -101.1412
23.4873 | -101.1362
23.491 | -101.1329
23.4948 | -101.13
23.4985 | -101.126
23.5023 | -101.1195
23.506 | -101.1129
23.5098 | -101.1066
23.5135 | -101.1005
23.5173 | -101.0945
23.521 | -101.0876
23.5248 | -101.0813
23.5285 | -101.0764
23.5323 | -101.0698
23.536 | -101.0621
23.5398 | -101.0548
23.5435 | -101.0498
23.5473 | -101.0434
23.551 | -101.0354
23.5548 | -101.0293
23.5585 | -101.0232
23.5623 | -101.0186
23.566 | -101.0125
23.5698 | -101.0067
23.5735 | -101.0025
23.5773 | -100.9956
23.581 | -100.9896
23.5848 | -100.9864
23.5885 | -100.9843
23.5923 | -100.981
23.596 | -100.9771
23.5998 | -100.9752
23.6035 | -100.9742
23.6073 | -100.9729
23.611 | -100.9716
23.6148 | -100.9718
23.6185 | -100.9741
23.6223 | -100.9752
23.626 | -100.9766
23.6298 | -100.9805
23.6336 | -100.9838
23.6373 | -100.9887
23.6411 | -100.9956
23.6448 | -101.0008
23.6486 | -101.0051
23.6523 | -101.0086
23.6561 | -101.0117
23.6598 | -101.0167
23.6636 | -101.0213
23.6673 | -101.0247
23.6711 | -101.0308
23.6748 | -101.036
23.6786 | -101.0398
23.6823 | -101.0473
23.6861 | -101.0573
23.6898 | -101.0671
23.6936 | -101.0749
23.6973 | -101.0774
23.7011 | -101.0797
23.7048 | -101.0865
23.7086 | -101.0947
23.7123 | -101.1041
23.7161 | -101.1113
23.7198 | -101.1171
23.7236 | -101.1244
23.7273 | -101.1344
23.7311 | -101.1446
23.7348 | -101.1543
23.7386 | -101.1656
23.7423 | -101.1735
23.7461 | -101.1792
23.7498 | -101.1863
23.7536 | -101.1924
23.7573 | -101.2012
23.7611 | -101.2098
23.7648 | -101.2178
23.7686 | -101.225
23.7723 | -101.2316
23.7761 | -101.2397
23.7798 | -101.2504
23.7836 | -101.2591
23.7873 | -101.2637
23.7911 | -101.2682
23.7948 | -101.2723
23.7986 | -101.2752
23.8023 | -101.2802
23.8061 | -101.2839
23.8098 | -101.2873
23.8136 | -101.2877
23.8173 | -101.2888
23.8211 | -101.2905
23.8248 | -101.291
23.8286 | -101.2904
23.8323 | -101.29
23.8361 | -101.2924
23.8398 | -101.2966
23.8436 | -101.3007
23.8473 | -101.3058
23.8511 | -101.3088
23.8548 | -101.3118
23.8586 | -101.3113
23.8623 | -101.3118
23.8661 | -101.3139
23.8698 | -101.3175
23.8736 | -101.3226
23.8773 | -101.3257
23.8811 | -101.327
23.8849 | -101.3305
23.8886 | -101.3357
23.8924 | -101.3399
23.8961 | -101.3414
23.8999 | -101.3418
23.9036 | -101.3407
23.9074 | -101.3417
23.9111 | -101.3446
23.9149 | -101.3478
23.9186 | -101.3513
23.9224 | -101.3533
23.9261 | -101.3535
23.9299 | -101.3529
23.9336 | -101.3529
23.9374 | -101.3544
23.9411 | -101.355
23.9449 | -101.3546
23.9486 | -101.355
23.9524 | -101.3535
23.9561 | -101.3483
23.9599 | -101.345
23.9636 | -101.344
23.9674 | -101.3395
23.9711 | -101.3352
23.9749 | -101.3301
23.9786 | -101.3216
23.9824 | -101.3138
23.9861 | -101.3062
23.9899 | -101.2984
23.9936 | -101.2907
23.9974 | -101.2802
24.0011 | -101.2712
24.0049 | -101.2629
24.0086 | -101.2524
24.0124 | -101.2422
24.0161 | -101.2336
24.0199 | -101.2234
24.0236 | -101.2128
24.0274 | -101.2013
24.0311 | -101.1879
24.0349 | -101.1796
24.0386 | -101.1698
24.0424 | -101.1561
24.0461 | -101.1413
24.0499 | -101.1253
24.0536 | -101.1066
24.0574 | -101.0883
24.0611 | -101.0724
24.0649 | -101.0566
24.0686 | -101.0409
24.0724 | -101.0267
24.0761 | -101.0122
24.0799 | -100.9991
24.0836 | -100.9875
24.0874 | -100.9755
24.0911 | -100.9636
24.0949 | -100.9485
24.0986 | -100.9368
24.1024 | -100.9267
24.1061 | -100.9167
24.1099 | -100.9079
24.1136 | -100.8973
24.1174 | -100.8846
24.1212 | -100.8738
24.1249 | -100.8636
24.1287 | -100.8545
24.1324 | -100.8454
24.1362 | -100.8379
24.1399 | -100.8312
24.1437 | -100.8247
24.1474 | -100.8174
24.1512 | -100.8124
24.1549 | -100.8079
24.1587 | -100.8012
24.1624 | -100.7969
24.1662 | -100.7933
24.1699 | -100.7892
24.1737 | -100.7848
24.1774 | -100.7804
24.1812 | -100.7786
24.1849 | -100.7763
24.1887 | -100.7743
24.1924 | -100.775
24.1962 | -100.775
24.1999 | -100.7743
24.2037 | -100.7741
24.2074 | -100.7733
24.2112 | -100.771
24.2149 | -100.7702
24.2187 | -100.7723
24.2224 | -100.774
24.2262 | -100.7751
24.2299 | -100.7773
24.2337 | -100.7789
24.2374 | -100.7832
24.2412 | -100.7868
24.2449 | -100.7903
24.2487 | -100.7922
24.2524 | -100.7931
24.2562 | -100.7919
24.2599 | -100.7825
24.2637 | -100.7681
24.2674 | -100.7564
24.2712 | -100.7525
24.2749 | -100.751
24.2787 | -100.7477
24.2824 | -100.7446
24.2862 | -100.7422
24.2899 | -100.7386
24.2937 | -100.7342
24.2974 | -100.7329
24.3012 | -100.7287
24.3049 | -100.7234
24.3087 | -100.7174
24.3124 | -100.7112
24.3162 | -100.7043
24.3199 | -100.6954
24.3237 | -100.6853
24.3274 | -100.6762
24.3312 | -100.6687
24.3349 | -100.6602
24.3387 | -100.651
24.3424 | -100.6397
24.3462 | -100.629
24.3499 | -100.6165
24.3537 | -100.6043
24.3574 | -100.5924
24.3612 | -100.5796
24.3649 | -100.5672
24.3687 | -100.5551
24.3725 | -100.5435
24.3762 | -100.5336
24.38 | -100.5228
24.3837 | -100.514
24.3875 | -100.5042
24.3912 | -100.4963
24.395 | -100.4925
24.3987 | -100.4906
24.4025 | -100.4894
24.4062 | -100.4861
24.41 | -100.4824
24.4137 | -100.4796
24.4175 | -100.477
24.4212 | -100.4774
24.425 | -100.4809
24.4287 | -100.4855
24.4325 | -100.4897
24.4362 | -100.4942
24.44 | -100.5022
24.4437 | -100.5097
24.4475 | -100.5186
24.4512 | -100.529
24.455 | -100.5421
24.4587 | -100.5592
24.4625 | -100.5788
24.4662 | -100.598
24.47 | -100.6195
24.4737 | -100.6407
24.4775 | -100.6635
24.4812 | -100.6862
24.485 | -100.7061
24.4887 | -100.727
24.4925 | -100.7509
24.4962 | -100.774
24.5 | -100.7987
24.5037 | -100.8243
24.5075 | -100.8476
24.5112 | -100.8721
24.515 | -100.8979
24.5187 | -100.925
24.5225 | -100.9505
24.5262 | -100.9731
24.53 | -100.997
24.5337 | -101.0204
24.5375 | -101.0422
24.5412 | -101.0664
24.545 | -101.0939
24.5487 | -101.1188
24.5525 | -101.1417
24.5562 | -101.1658
24.56 | -101.1891
24.5637 | -101.2117
24.5675 | -101.2332
24.5712 | -101.2546
24.575 | -101.2728
24.5787 | -101.2846
24.5825 | -101.2817
24.5862 | -101.2681
24.59 | -101.2476
24.5937 | -101.2364
24.5975 | -101.2322
24.6012 | -101.2334
24.605 | -101.2398
24.6087 | -101.247
24.6125 | -101.2453
24.6163 | -101.2297
24.62 | -101.2088
24.6238 | -101.1857
24.6275 | -101.1619
24.6313 | -101.138
24.635 | -101.1104
24.6388 | -101.0821
24.6425 | -101.0518
24.6463 | -101.0192
24.65 | -100.9845
24.6538 | -100.9488
24.6575 | -100.9112
24.6613 | -100.873
24.665 | -100.8368
24.6688 | -100.7984
24.6725 | -100.7568
24.6763 | -100.7121
24.68 | -100.6676
24.6838 | -100.6235
24.6875 | -100.5798
24.6913 | -100.5363
24.695 | -100.4937
24.6988 | -100.4522
24.7025 | -100.4116
24.7063 | -100.3706
24.71 | -100.3313
24.7138 | -100.2951
24.7175 | -100.2607
24.7213 | -100.2282
24.725 | -100.1989
24.7288 | -100.1722
24.7325 | -100.1477
24.7363 | -100.1281
24.74 | -100.1113
24.7438 | -100.0958
24.7475 | -100.0838
24.7513 | -100.0745
24.755 | -100.0686
24.7588 | -100.067
24.7625 | -100.0699
24.7663 | -100.077
24.77 | -100.0872
24.7738 | -100.1012
24.7775 | -100.1196
24.7813 | -100.1421
24.785 | -100.1697
24.7888 | -100.1999
24.7925 | -100.2334
24.7963 | -100.2698
24.8 | -100.3102
24.8038 | -100.3524
24.8075 | -100.3947
24.8113 | -100.4394
24.815 | -100.4867
24.8188 | -100.5349
24.8225 | -100.5843
24.8263 | -100.6363
24.83 | -100.6915
24.8338 | -100.7484
24.8375 | -100.8064
24.8413 | -100.8619
24.845 | -100.9184
24.8488 | -100.976
24.8525 | -101.0307
24.8563 | -101.0848
24.8601 | -101.1389
24.8638 | -101.1918
24.8676 | -101.2409
24.8713 | -101.2866
24.8751 | -101.3317
24.8788 | -101.3732
24.8826 | -101.4126
24.8863 | -101.4521
24.8901 | -101.4646
24.8938 | -101.4958
24.8976 | -101.5242
24.9013 | -101.5522
24.9051 | -101.5788
24.9088 | -101.602
24.9126 | -101.6216
24.9163 | -101.638
24.9201 | -101.654
24.9238 | -101.6671
24.9276 | -101.6785
24.9313 | -101.695
24.9351 | -101.7223
24.9388 | -101.754
24.9426 | -101.7874
24.9463 | -101.8014
24.9501 | -101.8004
24.9538 | -101.7935
24.9576 | -101.7871
24.9613 | -101.7809
24.9651 | -101.7766
24.9688 | -101.7747
24.9726 | -101.7736
24.9763 | -101.7726
24.9801 | -101.771
24.9838 | -101.7681
24.9876 | -101.7677
24.9913 | -101.7675
24.9951 | -101.7679
24.9988 | -101.7694
25.0026 | -101.7723
25.0063 | -101.777
25.0101 | -101.7836
25.0138 | -101.7924
25.0176 | -101.8004
25.0213 | -101.8093
25.0251 | -101.819
25.0288 | -101.832
25.0326 | -101.8467
25.0363 | -101.8632
25.0401 | -101.8805
25.0438 | -101.8984
25.0476 | -101.9167
25.0513 | -101.9396
25.0551 | -101.9639
25.0588 | -101.9887
25.0626 | -102.0135
25.0663 | -102.0378
25.0701 | -102.0631
25.0738 | -102.0906
25.0776 | -102.119
25.0813 | -102.1474
25.0851 | -102.1738
25.0888 | -102.2007
25.0926 | -102.229
25.0963 | -102.2577
25.1001 | -102.2897
25.1038 | -102.3223
25.1076 | -102.3562
25.1114 | -102.3895
25.1151 | -102.4205
25.1189 | -102.4502
25.1226 | -102.4799
25.1264 | -102.5087
25.1301 | -102.5363
25.1339 | -102.5635
25.1376 | -102.5881
25.1414 | -102.6117
25.1451 | -102.6344
25.1489 | -102.6583
25.1526 | -102.6809
25.1564 | -102.7013
25.1601 | -102.7253
25.1639 | -102.7452
25.1676 | -102.7634
25.1714 | -102.782
25.1751 | -102.8002
25.1789 | -102.8156
25.1826 | -102.8271
25.1864 | -102.8397
25.1901 | -102.8494
25.1939 | -102.8603
25.1976 | -102.8694
25.2014 | -102.8749
25.2051 | -102.8806
25.2089 | -102.8839
25.2126 | -102.8863
25.2164 | -102.8894
25.2201 | -102.8941
25.2239 | -102.9015
25.2276 | -102.9066
25.2314 | -102.9118
25.2351 | -102.9161
25.2389 | -102.919
25.2426 | -102.9226
25.2464 | -102.9259
25.2501 | -102.9308
25.2539 | -102.9364
25.2576 | -102.9407
25.2614 | -102.9441
25.2651 | -102.9474
25.2689 | -102.9518
25.2726 | -102.9583
25.2764 | -102.9651
25.2801 | -102.9746
25.2839 | -102.9837
25.2876 | -102.9911
25.2914 | -102.9978
25.2951 | -103.0049
25.2989 | -103.014
25.3026 | -103.0252
25.3064 | -103.0366
25.3101 | -103.0484
25.3139 | -103.0611
25.3176 | -103.0744
25.3214 | -103.0903
25.3251 | -103.1051
25.3289 | -103.1202
25.3326 | -103.1362
25.3364 | -103.1528
25.3401 | -103.171
25.3439 | -103.1893
25.3476 | -103.2079
25.3514 | -103.2258
25.3551 | -103.2449
25.3589 | -103.2623
25.3627 | -103.2812
25.3664 | -103.3001
25.3702 | -103.3191
25.3739 | -103.3393
25.3777 | -103.3603
25.3814 | -103.3809
25.3852 | -103.3983
25.3889 | -103.4143
25.3927 | -103.431
25.3964 | -103.4482
25.4002 | -103.4655
25.4039 | -103.4814
25.4077 | -103.4932
25.4114 | -103.5026
25.4152 | -103.5145
25.4189 | -103.5275
25.4227 | -103.5392
25.4264 | -103.5498
25.4302 | -103.5503
25.4339 | -103.5595
25.4377 | -103.5669
25.4414 | -103.5741
25.4452 | -103.5835
25.4489 | -103.5913
25.4527 | -103.5964
25.4564 | -103.6005
25.4602 | -103.6017
25.4639 | -103.6029
25.4677 | -103.6046
25.4714 | -103.6063
25.4752 | -103.6075
25.4789 | -103.6092
25.4827 | -103.6094
25.4864 | -103.6091
25.4902 | -103.6065
25.4939 | -103.6057
25.4977 | -103.6041
25.5014 | -103.6035
25.5052 | -103.6028
25.5089 | -103.5984
25.5127 | -103.5942
25.5164 | -103.5915
25.5202 | -103.5877
25.5239 | -103.5858
25.5277 | -103.5837
25.5314 | -103.5778
25.5352 | -103.5689
25.5389 | -103.5595
25.5427 | -103.5529
25.5464 | -103.5446
25.5502 | -103.5369
25.5539 | -103.5301
25.5577 | -103.5234
25.5614 | -103.5165
25.5652 | -103.5103
25.5689 | -103.5034
25.5727 | -103.4953
25.5764 | -103.4863
25.5802 | -103.4781
25.5839 | -103.4699
25.5877 | -103.4632
25.5914 | -103.4568
25.5952 | -103.4496
25.5989 | -103.4417
25.6027 | -103.4335
25.6065 | -103.4276
25.6102 | -103.4235
25.614 | -103.4194
25.6177 | -103.4141
25.6215 | -103.4098
25.6252 | -103.4073
25.629 | -103.4048
25.6327 | -103.4032
25.6365 | -103.4017
25.6402 | -103.3985
25.644 | -103.3983
25.6477 | -103.3993
25.6515 | -103.4017
25.6552 | -103.4055
25.659 | -103.4077
25.6627 | -103.4103
25.6665 | -103.4148
25.6702 | -103.4211
25.674 | -103.4278
25.6777 | -103.4345
25.6815 | -103.4432
25.6852 | -103.4546
25.689 | -103.4666
25.6927 | -103.4764
25.6965 | -103.486
25.7002 | -103.4959
25.704 | -103.5082
25.7077 | -103.5231
25.7115 | -103.5395
25.7152 | -103.5534
25.719 | -103.566
25.7227 | -103.5796
25.7265 | -103.5943
25.7302 | -103.6091
25.734 | -103.6249
25.7377 | -103.6405
25.7415 | -103.6568
25.7452 | -103.6726
25.749 | -103.6893
25.7527 | -103.7053
25.7565 | -103.7211
25.7602 | -103.7356
25.764 | -103.7494
25.7677 | -103.7634
25.7715 | -103.7782
25.7752 | -103.7909
25.779 | -103.8017
25.7827 | -103.8115
25.7865 | -103.8191
25.7902 | -103.8262
25.794 | -103.8328
25.7977 | -103.8372
25.8015 | -103.8425
25.8052 | -103.8475
25.809 | -103.8503
25.8127 | -103.8521
25.8165 | -103.8545
25.8202 | -103.8577
25.824 | -103.8605
25.8277 | -103.8621
25.8315 | -103.8608
25.8352 | -103.858
25.839 | -103.8555
25.8427 | -103.8533
25.8465 | -103.8506
25.8503 | -103.8476
25.854 | -103.8473
25.8578 | -103.8465
25.8615 | -103.8442
25.8653 | -103.8405
25.869 | -103.8363
25.8728 | -103.8329
25.8765 | -103.831
25.8803 | -103.8275
25.884 | -103.8242
25.8878 | -103.8193
25.8915 | -103.8132
25.8953 | -103.8114
25.899 | -103.8125
25.9028 | -103.8132
25.9065 | -103.8113
25.9103 | -103.8099
25.914 | -103.8077
25.9178 | -103.8047
25.9215 | -103.8026
25.9253 | -103.802
25.929 | -103.8016
25.9328 | -103.8012
25.9365 | -103.7986
25.9403 | -103.7963
25.944 | -103.7942
25.9478 | -103.792
25.9515 | -103.7885
25.9553 | -103.7869
25.959 | -103.7848
25.9628 | -103.7822
25.9665 | -103.7811
25.9703 | -103.7774
25.974 | -103.7757
25.9778 | -103.7729
25.9815 | -103.7679
25.9853 | -103.7585
25.989 | -103.7498
25.9928 | -103.7443
25.9965 | -103.7403
26.0003 | -103.7338
26.004 | -103.7268
26.0078 | -103.7195
26.0115 | -103.712
26.0153 | -103.7042
26.019 | -103.6956
26.0228 | -103.6872
26.0265 | -103.678
26.0303 | -103.6676
26.034 | -103.6582
26.0378 | -103.6463
26.0415 | -103.6347
26.0453 | -103.6225
26.049 | -103.6077
26.0528 | -103.5925
26.0565 | -103.578
26.0603 | -103.5666
26.064 | -103.5569
26.0678 | -103.547
26.0715 | -103.5354
26.0753 | -103.5222
26.079 | -103.5102
26.0828 | -103.5009
26.0865 | -103.493
26.0903 | -103.4827
26.094 | -103.4738
26.0978 | -103.4653
26.1016 | -103.4565
26.1053 | -103.4492
26.1091 | -103.4409
26.1128 | -103.4339
26.1166 | -103.4281
26.1203 | -103.4243
26.1241 | -103.4202
26.1278 | -103.4178
26.1316 | -103.4187
26.1353 | -103.4197
26.1391 | -103.4215
26.1428 | -103.4233
26.1466 | -103.4259
26.1503 | -103.4298
26.1541 | -103.4366
26.1578 | -103.4445
26.1616 | -103.4523
26.1653 | -103.4624
26.1691 | -103.4718
26.1728 | -103.4804
26.1766 | -103.4901
26.1803 | -103.4994
26.1841 | -103.5088
26.1878 | -103.5206
26.1916 | -103.5324
26.1953 | -103.5454
26.1991 | -103.5612
26.2028 | -103.5773
26.2066 | -103.5939
26.2103 | -103.6088
26.2141 | -103.6257
26.2178 | -103.6433
26.2216 | -103.6604
26.2253 | -103.6761
26.2291 | -103.6941
26.2328 | -103.7119
26.2366 | -103.7278
26.2403 | -103.7432
26.2441 | -103.7575
26.2478 | -103.7726
26.2516 | -103.7863
26.2553 | -103.8016
26.2591 | -103.8175
26.2628 | -103.8307
26.2666 | -103.8415
26.2703 | -103.851
26.2741 | -103.8605
26.2778 | -103.87
26.2816 | -103.8803
26.2853 | -103.8911
26.2891 | -103.9004
26.2928 | -103.9067
26.2966 | -103.9125
26.3003 | -103.9199
26.3041 | -103.9284
26.3078 | -103.9343
26.3116 | -103.9382
26.3153 | -103.9411
26.3191 | -103.944
26.3228 | -103.945
26.3266 | -103.9452
26.3303 | -103.9434
26.3341 | -103.9422
26.3378 | -103.9403
26.3416 | -103.9362
26.3454 | -103.9325
26.3491 | -103.9291
26.3529 | -103.9255
26.3566 | -103.9251
26.3604 | -103.9213
26.3641 | -103.9151
26.3679 | -103.9084
26.3716 | -103.9013
26.3754 | -103.8922
26.3791 | -103.8846
26.3829 | -103.8763
26.3866 | -103.8649
26.3904 | -103.8562
26.3941 | -103.8494
26.3979 | -103.8416
26.4016 | -103.8326
26.4054 | -103.8215
26.4091 | -103.8111
26.4129 | -103.7998
26.4166 | -103.7899
26.4204 | -103.7827
26.4241 | -103.7764
26.4279 | -103.7699
26.4316 | -103.7622
26.4354 | -103.7547
26.4391 | -103.7478
26.4429 | -103.7415
26.4466 | -103.7349
26.4504 | -103.727
26.4541 | -103.7199
26.4579 | -103.7143
26.4616 | -103.709
26.4654 | -103.7025
26.4691 | -103.6955
26.4729 | -103.6896
26.4766 | -103.6855
26.4804 | -103.682
26.4841 | -103.6778
26.4879 | -103.6718
26.4916 | -103.6639
26.4954 | -103.6554
26.4991 | -103.6476
26.5029 | -103.641
26.5066 | -103.6345
26.5104 | -103.628
26.5141 | -103.6106
26.5179 | -103.5996
26.5216 | -103.5916
26.5254 | -103.584
26.5291 | -103.5746
26.5329 | -103.5644
26.5366 | -103.5521
26.5404 | -103.5382
26.5441 | -103.5261
26.5479 | -103.5152
26.5516 | -103.5066
26.5554 | -103.4977
26.5591 | -103.4881
26.5629 | -103.4754
26.5666 | -103.463
26.5704 | -103.4503
26.5741 | -103.4374
26.5779 | -103.4247
26.5816 | -103.4123
26.5854 | -103.4005
26.5892 | -103.3858
26.5929 | -103.3705
26.5967 | -103.3558
26.6004 | -103.3422
26.6042 | -103.3264
26.6079 | -103.3109
26.6117 | -103.2983
26.6154 | -103.2851
26.6192 | -103.2687
26.6229 | -103.2537
26.6267 | -103.2403
26.6304 | -103.2252
26.6342 | -103.2068
26.6379 | -103.1871
26.6417 | -103.1691
26.6454 | -103.1514
26.6492 | -103.1352
26.6529 | -103.1192
26.6567 | -103.1003
26.6604 | -103.0801
26.6642 | -103.0635
26.6679 | -103.0497
26.6717 | -103.0342
26.6754 | -103.0154
26.6792 | -102.9978
26.6829 | -102.9821
26.6867 | -102.9664
26.6904 | -102.9518
26.6942 | -102.9379
26.6979 | -102.9234
26.7017 | -102.9089
26.7054 | -102.8959
26.7092 | -102.8802
26.7129 | -102.8656
26.7167 | -102.8525
26.7204 | -102.8402
26.7242 | -102.8267
26.7279 | -102.8122
26.7317 | -102.8002
26.7354 | -102.7867
26.7392 | -102.7727
26.7429 | -102.7598
26.7467 | -102.746
26.7504 | -102.7315
26.7542 | -102.7169
26.7579 | -102.7025
26.7617 | -102.692
26.7654 | -102.6794
26.7692 | -102.6659
26.7729 | -102.6529
26.7767 | -102.6384
26.7804 | -102.6246
26.7842 | -102.6117
26.7879 | -102.6003
26.7917 | -102.589
26.7954 | -102.576
26.7992 | -102.5589
26.8029 | -102.5423
26.8067 | -102.5265
26.8104 | -102.5096
26.8142 | -102.4949
26.8179 | -102.4807
26.8217 | -102.4659
26.8254 | -102.4496
26.8292 | -102.4327
26.833 | -102.4173
26.8367 | -102.403
26.8405 | -102.3885
26.8442 | -102.3744
26.848 | -102.3594
26.8517 | -102.3426
26.8555 | -102.3266
26.8592 | -102.3123
26.863 | -102.2988
26.8667 | -102.285
26.8705 | -102.2728
26.8742 | -102.2595
26.878 | -102.2443
26.8817 | -102.2313
26.8855 | -102.219
26.8892 | -102.2073
26.893 | -102.1962
26.8967 | -102.182
26.9005 | -102.1676
26.9042 | -102.153
26.908 | -102.1412
26.9117 | -102.1293
26.9155 | -102.1171
26.9192 | -102.1034
26.923 | -102.0886
26.9267 | -102.0747
26.9305 | -102.0602
26.9342 | -102.0468
26.938 | -102.0327
26.9417 | -102.0207
26.9455 | -102.0073
26.9492 | -101.9941
26.953 | -101.9813
26.9567 | -101.9689
26.9605 | -101.9566
26.9642 | -101.9433
26.968 | -101.9302
26.9717 | -101.9173
26.9755 | -101.9053
26.9792 | -101.8932
26.983 | -101.879
26.9867 | -101.8677
26.9905 | -101.8542
26.9942 | -101.8394
26.998 | -101.8245
27.0017 | -101.8105
27.0055 | -101.7986
27.0092 | -101.7881
27.013 | -101.7795
27.0167 | -101.7724
27.0205 | -101.7622
27.0242 | -101.7483
27.028 | -101.7364
27.0317 | -101.7259
27.0355 | -101.7159
27.0392 | -101.7067
27.043 | -101.6949
27.0467 | -101.6811
27.0505 | -101.6676
27.0542 | -101.6517
27.058 | -101.6428
27.0617 | -101.631
27.0655 | -101.6192
27.0692 | -101.6069
27.073 | -101.5947
27.0768 | -101.5821
27.0805 | -101.5665
27.0843 | -101.5468
27.088 | -101.5277
27.0918 | -101.511
27.0955 | -101.4935
27.0993 | -101.4771
27.103 | -101.4577
27.1068 | -101.4391
27.1105 | -101.4197
27.1143 | -101.4
27.118 | -101.3831
27.1218 | -101.3657
27.1255 | -101.3472
27.1293 | -101.331
27.133 | -101.3157
27.1368 | -101.298
27.1405 | -101.2771
27.1443 | -101.2546
27.148 | -101.2346
27.1518 | -101.2152
27.1555 | -101.1969
27.1593 | -101.1792
27.163 | -101.1606
27.1668 | -101.1422
27.1705 | -101.1229
27.1743 | -101.1044
27.178 | -101.0858
27.1818 | -101.0701
27.1855 | -101.055
27.1893 | -101.041
27.193 | -101.0282
27.1968 | -101.0133
27.2005 | -100.9997
27.2043 | -100.9877
27.208 | -100.9789
27.2118 | -100.9697
27.2155 | -100.9587
27.2193 | -100.9485
27.223 | -100.94
27.2268 | -100.9343
27.2305 | -100.9298
27.2343 | -100.9266
27.238 | -100.9257
27.2418 | -100.9247
27.2455 | -100.9236
27.2493 | -100.9244
27.253 | -100.9251
27.2568 | -100.928
27.2605 | -100.9324
27.2643 | -100.9355
27.268 | -100.9389
27.2718 | -100.9458
27.2755 | -100.9552
27.2793 | -100.9664
27.283 | -100.9769
27.2868 | -100.9888
27.2905 | -100.9998
27.2943 | -101.0131
27.298 | -101.0298
27.3018 | -101.0481
27.3055 | -101.0674
27.3093 | -101.0865
27.313 | -101.1045
27.3168 | -101.1237
27.3205 | -101.1452
27.3243 | -101.1666
27.3281 | -101.1898
27.3318 | -101.213
27.3356 | -101.2358
27.3393 | -101.2607
27.3431 | -101.2864
27.3468 | -101.3134
27.3506 | -101.3409
27.3543 | -101.368
27.3581 | -101.3938
27.3618 | -101.419
27.3656 | -101.4469
27.3693 | -101.4719
27.3731 | -101.4983
27.3768 | -101.5268
27.3806 | -101.5543
27.3843 | -101.5796
27.3881 | -101.6015
27.3918 | -101.6238
27.3956 | -101.6461
27.3993 | -101.6667
27.4031 | -101.6877
27.4068 | -101.7116
27.4106 | -101.7308
27.4143 | -101.7497
27.4181 | -101.7698
27.4218 | -101.786
27.4256 | -101.8019
27.4293 | -101.8175
27.4331 | -101.8339
27.4368 | -101.8512
27.4406 | -101.8659
27.4443 | -101.878
27.4481 | -101.8916
27.4518 | -101.9039
27.4556 | -101.9162
27.4593 | -101.9266
27.4631 | -101.9353
27.4668 | -101.9481
27.4706 | -101.9659
27.4743 | -101.9807
27.4781 | -101.9881
27.4818 | -101.9943
27.4856 | -102.0008
27.4893 | -102.0096
27.4931 | -102.0165
27.4968 | -102.0254
27.5006 | -102.0311
27.5043 | -102.0326
27.5081 | -102.0343
27.5118 | -102.0379
27.5156 | -102.0398
27.5193 | -102.0423
27.5231 | -102.0475
27.5268 | -102.0553
27.5306 | -102.0629
27.5343 | -102.0659
27.5381 | -102.0674
27.5418 | -102.0672
27.5456 | -102.0672
27.5493 | -102.0652
27.5531 | -102.0604
27.5568 | -102.0557
27.5606 | -102.0522
27.5643 | -102.0511
27.5681 | -102.0495
27.5719 | -102.0482
27.5756 | -102.0447
27.5794 | -102.0397
27.5831 | -102.0363
27.5869 | -102.0336
27.5906 | -102.0306
27.5944 | -102.0202
27.5981 | -102.0132
27.6019 | -102.0075
27.6056 | -102.0048
27.6094 | -102.0005
27.6131 | -101.9943
27.6169 | -101.9869
27.6206 | -101.98
27.6244 | -101.9721
27.6281 | -101.9623
27.6319 | -101.9529
27.6356 | -101.9442
27.6394 | -101.9352
27.6431 | -101.9268
27.6469 | -101.9178
27.6506 | -101.9086
27.6544 | -101.9013
27.6581 | -101.8942
27.6619 | -101.8865
27.6656 | -101.8775
27.6694 | -101.8681
27.6731 | -101.8593
27.6769 | -101.8498
27.6806 | -101.8418
27.6844 | -101.8353
27.6881 | -101.827
27.6919 | -101.8177
27.6956 | -101.8085
27.6994 | -101.8005
27.7031 | -101.795
27.7069 | -101.788
27.7106 | -101.7827
27.7144 | -101.7777
27.7181 | -101.773
27.7219 | -101.7707
27.7256 | -101.7706
27.7294 | -101.7705
27.7331 | -101.7685
27.7369 | -101.7664
27.7406 | -101.7666
27.7444 | -101.7673
27.7481 | -101.7692
27.7519 | -101.7726
27.7556 | -101.7762
27.7594 | -101.7797
27.7631 | -101.7824
27.7669 | -101.7844
27.7706 | -101.79
27.7744 | -101.7963
27.7781 | -101.8031
27.7819 | -101.8092
27.7856 | -101.8156
27.7894 | -101.8233
27.7931 | -101.8312
27.7969 | -101.8386
27.8006 | -101.8469
27.8044 | -101.8553
27.8081 | -101.8623
27.8119 | -101.8719
27.8157 | -101.8816
27.8194 | -101.8918
27.8232 | -101.9016
27.8269 | -101.912
27.8307 | -101.9217
27.8344 | -101.931
27.8382 | -101.9399
27.8419 | -101.9476
27.8457 | -101.9549
27.8494 | -101.9635
27.8532 | -101.9724
27.8569 | -101.9792
27.8607 | -101.9858
27.8644 | -101.9922
27.8682 | -102.0013
27.8719 | -102.0104
27.8757 | -102.0174
27.8794 | -102.0242
27.8832 | -102.0309
27.8869 | -102.0349
27.8907 | -102.0388
27.8944 | -102.046
27.8982 | -102.0523
27.9019 | -102.0568
27.9057 | -102.06
27.9094 | -102.0609
27.9132 | -102.0609
27.9169 | -102.0605
27.9207 | -102.0588
27.9244 | -102.0594
27.9282 | -102.0597
27.9319 | -102.0601
27.9357 | -102.0626
27.9394 | -102.0658
27.9432 | -102.0697
27.9469 | -102.0731
27.9507 | -102.0737
27.9544 | -102.0701
27.9582 | -102.0648
27.9619 | -102.0576
27.9657 | -102.0522
27.9694 | -102.0489
27.9732 | -102.0451
27.9769 | -102.0371
27.9807 | -102.0294
27.9844 | -102.0239
27.9882 | -102.0178
27.9919 | -102.0124
27.9957 | -102.0052
27.9994 | -101.9956
28.0032 | -101.9852