## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 253 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 272
Hood | 02-92B/P | 199
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1014.1 | hPa |  | PRC PRO 01
Humidity | 56.4 | %/Hr
Internal temperature of calibration bench | 23.15 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1017 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.06 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.98 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3094 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 20/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 25/06/2018
Technical Manager | Anthony HUE | Approval |  | 28/06/2018
CEO | Robert GLEMOT | Approval |  | 28/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4587 | 57.555
0.012209 | -35.9898 | 38.515
0.018313 | -34.9427 | 28.5444
0.024417 | -34.492 | 22.1443
0.030521 | -34.3816 | 18.3475
0.036626 | -34.2386 | 15.2435
0.04273 | -34.1556 | 12.9325
0.048834 | -34.1035 | 11.3446
0.054938 | -34.0342 | 10.1394
0.061043 | -33.9608 | 9.01
0.067147 | -33.9505 | 8.142
0.073251 | -33.951 | 7.3879
0.079355 | -33.9679 | 6.8274
0.08546 | -33.985 | 6.3684
0.091564 | -33.9814 | 5.9908
0.097668 | -33.9475 | 5.7287
0.10377 | -33.9004 | 5.3797
0.10988 | -33.8699 | 5.016
0.11598 | -33.861 | 4.6562
0.12209 | -33.8776 | 4.3764
0.12819 | -33.9099 | 4.1428
0.13429 | -33.9412 | 3.9034
0.1404 | -33.9581 | 3.6364
0.1465 | -33.9738 | 3.3833
0.15261 | -33.9811 | 3.148
0.15871 | -33.9781 | 2.9527
0.16482 | -33.9642 | 2.7911
0.17092 | -33.9396 | 2.7222
0.17702 | -33.9207 | 2.6203
0.18313 | -33.9034 | 2.5343
0.18923 | -33.8909 | 2.4278
0.19534 | -33.8841 | 2.3258
0.20144 | -33.8833 | 2.2155
0.20754 | -33.8863 | 2.1444
0.21365 | -33.8888 | 2.0716
0.21975 | -33.8933 | 2.0079
0.22586 | -33.9013 | 1.9431
0.23196 | -33.9098 | 1.8698
0.23807 | -33.9182 | 1.7908
0.24417 | -33.9201 | 1.6964
0.25027 | -33.9231 | 1.6115
0.25638 | -33.9228 | 1.519
0.26248 | -33.9207 | 1.4323
0.26859 | -33.9182 | 1.3496
0.27469 | -33.9174 | 1.2641
0.2808 | -33.9233 | 1.2012
0.2869 | -33.9278 | 1.1339
0.293 | -33.9332 | 1.0776
0.29911 | -33.9395 | 1.0397
0.30521 | -33.9455 | 1.0036
0.31132 | -33.9496 | 0.98909
0.31742 | -33.9539 | 0.96659
0.32353 | -33.9574 | 0.94771
0.32963 | -33.9603 | 0.93163
0.33573 | -33.9622 | 0.91728
0.34184 | -33.9647 | 0.89972
0.34794 | -33.9678 | 0.86386
0.35405 | -33.971 | 0.83307
0.36015 | -33.9745 | 0.79898
0.36626 | -33.9767 | 0.7616
0.37236 | -33.9791 | 0.71204
0.37846 | -33.9811 | 0.65561
0.38457 | -33.98 | 0.59847
0.39067 | -33.9806 | 0.54943
0.39678 | -33.9788 | 0.49463
0.40288 | -33.9764 | 0.44316
0.40899 | -33.9725 | 0.39288
0.41509 | -33.968 | 0.34222
0.42119 | -33.963 | 0.29576
0.4273 | -33.9592 | 0.24789
0.4334 | -33.9558 | 0.19912
0.43951 | -33.9527 | 0.15097
0.44561 | -33.9511 | 0.1057
0.45172 | -33.9526 | 0.067825
0.45782 | -33.9523 | 0.026112
0.46392 | -33.9535 | -0.012166
0.47003 | -33.9551 | -0.047261
0.47613 | -33.9572 | -0.079079
0.48224 | -33.9594 | -0.10919
0.48834 | -33.962 | -0.12992
0.49445 | -33.9641 | -0.15407
0.50055 | -33.9668 | -0.176
0.50665 | -33.9687 | -0.19347
0.51276 | -33.97 | -0.21238
0.51886 | -33.9713 | -0.22875
0.52497 | -33.9704 | -0.23455
0.53107 | -33.971 | -0.24395
0.53717 | -33.9714 | -0.25434
0.54328 | -33.9713 | -0.2589
0.54938 | -33.9714 | -0.26645
0.55549 | -33.9709 | -0.27499
0.56159 | -33.9708 | -0.28621
0.5677 | -33.9704 | -0.29665
0.5738 | -33.9699 | -0.30979
0.5799 | -33.9699 | -0.32184
0.58601 | -33.9703 | -0.33872
0.59211 | -33.9713 | -0.36206
0.59822 | -33.9719 | -0.37922
0.60432 | -33.9729 | -0.40117
0.61043 | -33.9738 | -0.42309
0.61653 | -33.975 | -0.44516
0.62263 | -33.976 | -0.46888
0.62874 | -33.9774 | -0.48974
0.63484 | -33.9785 | -0.51152
0.64095 | -33.9798 | -0.53315
0.64705 | -33.9809 | -0.55373
0.65316 | -33.982 | -0.57363
0.65926 | -33.983 | -0.59632
0.66536 | -33.9826 | -0.61461
0.67147 | -33.9829 | -0.63117
0.67757 | -33.9835 | -0.64931
0.68368 | -33.9833 | -0.66912
0.68978 | -33.9835 | -0.68703
0.69589 | -33.9832 | -0.70509
0.70199 | -33.9825 | -0.72639
0.70809 | -33.982 | -0.74628
0.7142 | -33.9816 | -0.76512
0.7203 | -33.9808 | -0.78435
0.72641 | -33.98 | -0.80456
0.73251 | -33.9798 | -0.82492
0.73862 | -33.9799 | -0.84865
0.74472 | -33.9802 | -0.86756
0.75082 | -33.9798 | -0.88799
0.75693 | -33.9798 | -0.91095
0.76303 | -33.9798 | -0.9329
0.76914 | -33.9802 | -0.95958
0.77524 | -33.9803 | -0.98418
0.78135 | -33.9807 | -1.0072
0.78745 | -33.9812 | -1.0317
0.79355 | -33.9817 | -1.056
0.79966 | -33.9825 | -1.0816
0.80576 | -33.9836 | -1.1118
0.81187 | -33.9838 | -1.141
0.81797 | -33.9843 | -1.1714
0.82408 | -33.9849 | -1.1983
0.83018 | -33.9854 | -1.2292
0.83628 | -33.9861 | -1.2597
0.84239 | -33.9859 | -1.2875
0.84849 | -33.9859 | -1.3152
0.8546 | -33.9862 | -1.3426
0.8607 | -33.9862 | -1.3716
0.86681 | -33.9863 | -1.3997
0.87291 | -33.9858 | -1.4252
0.87901 | -33.9852 | -1.4472
0.88512 | -33.9851 | -1.47
0.89122 | -33.9849 | -1.493
0.89733 | -33.9846 | -1.5121
0.90343 | -33.9843 | -1.5331
0.90953 | -33.9845 | -1.5477
0.91564 | -33.9839 | -1.5633
0.92174 | -33.9835 | -1.5772
0.92785 | -33.983 | -1.5905
0.93395 | -33.983 | -1.6019
0.94006 | -33.983 | -1.6154
0.94616 | -33.9829 | -1.6197
0.95226 | -33.9829 | -1.6309
0.95837 | -33.9831 | -1.6423
0.96447 | -33.9835 | -1.6522
0.97058 | -33.9835 | -1.6618
0.97668 | -33.9833 | -1.6694
0.98279 | -33.9833 | -1.6725
0.98889 | -33.9836 | -1.6771
0.99499 | -33.9836 | -1.6852
1.0011 | -33.9839 | -1.6898
1.0072 | -33.9838 | -1.6932
1.0133 | -33.9838 | -1.6968
1.0194 | -33.9838 | -1.7018
1.0255 | -33.9835 | -1.7082
1.0316 | -33.9829 | -1.712
1.0377 | -33.9824 | -1.7174
1.0438 | -33.9817 | -1.7214
1.0499 | -33.9802 | -1.7273
1.056 | -33.9791 | -1.7316
1.0621 | -33.978 | -1.7369
1.0682 | -33.9775 | -1.7414
1.0743 | -33.9764 | -1.7487
1.0805 | -33.9758 | -1.755
1.0866 | -33.9743 | -1.7654
1.0927 | -33.9729 | -1.7723
1.0988 | -33.9722 | -1.7825
1.1049 | -33.9708 | -1.7896
1.111 | -33.9692 | -1.8002
1.1171 | -33.9679 | -1.8092
1.1232 | -33.9665 | -1.823
1.1293 | -33.9654 | -1.8346
1.1354 | -33.9645 | -1.8465
1.1415 | -33.9633 | -1.8602
1.1476 | -33.9619 | -1.8737
1.1537 | -33.9606 | -1.8866
1.1598 | -33.9604 | -1.9047
1.1659 | -33.9599 | -1.9164
1.172 | -33.9592 | -1.9294
1.1781 | -33.9585 | -1.9431
1.1842 | -33.958 | -1.9585
1.1903 | -33.9587 | -1.9736
1.1964 | -33.9586 | -1.9879
1.2025 | -33.9585 | -2.0029
1.2086 | -33.959 | -2.0185
1.2147 | -33.9592 | -2.0333
1.2209 | -33.9599 | -2.0467
1.227 | -33.9618 | -2.0617
1.2331 | -33.9628 | -2.0759
1.2392 | -33.9637 | -2.0916
1.2453 | -33.965 | -2.1036
1.2514 | -33.9663 | -2.1157
1.2575 | -33.9678 | -2.1297
1.2636 | -33.9702 | -2.1412
1.2697 | -33.9719 | -2.153
1.2758 | -33.9735 | -2.167
1.2819 | -33.9752 | -2.1806
1.288 | -33.9769 | -2.1944
1.2941 | -33.979 | -2.2051
1.3002 | -33.981 | -2.2135
1.3063 | -33.9827 | -2.2235
1.3124 | -33.9845 | -2.2335
1.3185 | -33.9865 | -2.2443
1.3246 | -33.9883 | -2.2551
1.3307 | -33.9904 | -2.2646
1.3368 | -33.9923 | -2.2722
1.3429 | -33.9939 | -2.281
1.349 | -33.9954 | -2.2895
1.3551 | -33.9971 | -2.3004
1.3613 | -33.9987 | -2.3083
1.3674 | -34 | -2.3167
1.3735 | -34.0013 | -2.3247
1.3796 | -34.0025 | -2.3321
1.3857 | -34.0036 | -2.339
1.3918 | -34.0046 | -2.3487
1.3979 | -34.0059 | -2.3572
1.404 | -34.0066 | -2.367
1.4101 | -34.0075 | -2.3759
1.4162 | -34.0084 | -2.3843
1.4223 | -34.0092 | -2.3945
1.4284 | -34.01 | -2.4027
1.4345 | -34.0106 | -2.4128
1.4406 | -34.0111 | -2.4237
1.4467 | -34.0118 | -2.4329
1.4528 | -34.0124 | -2.4419
1.4589 | -34.0132 | -2.4523
1.465 | -34.0135 | -2.4598
1.4711 | -34.0144 | -2.4686
1.4772 | -34.0149 | -2.4805
1.4833 | -34.0155 | -2.4925
1.4894 | -34.0162 | -2.5037
1.4955 | -34.0169 | -2.5151
1.5016 | -34.0176 | -2.5269
1.5078 | -34.018 | -2.5412
1.5139 | -34.0187 | -2.5535
1.52 | -34.0191 | -2.5665
1.5261 | -34.0196 | -2.578
1.5322 | -34.0197 | -2.5902
1.5383 | -34.02 | -2.6022
1.5444 | -34.0206 | -2.6153
1.5505 | -34.021 | -2.6283
1.5566 | -34.0215 | -2.6408
1.5627 | -34.0221 | -2.6539
1.5688 | -34.0229 | -2.6668
1.5749 | -34.0231 | -2.6789
1.581 | -34.0234 | -2.6953
1.5871 | -34.024 | -2.7107
1.5932 | -34.0244 | -2.7233
1.5993 | -34.0247 | -2.7398
1.6054 | -34.0251 | -2.7545
1.6115 | -34.0254 | -2.7712
1.6176 | -34.0257 | -2.7874
1.6237 | -34.026 | -2.8015
1.6298 | -34.0262 | -2.8167
1.6359 | -34.0266 | -2.831
1.642 | -34.0271 | -2.8453
1.6482 | -34.0274 | -2.8621
1.6543 | -34.0279 | -2.8787
1.6604 | -34.0283 | -2.8931
1.6665 | -34.0284 | -2.9079
1.6726 | -34.0287 | -2.9239
1.6787 | -34.0292 | -2.9407
1.6848 | -34.029 | -2.957
1.6909 | -34.0294 | -2.973
1.697 | -34.0296 | -2.9891
1.7031 | -34.0296 | -3.0046
1.7092 | -34.0296 | -3.0193
1.7153 | -34.0298 | -3.0353
1.7214 | -34.0294 | -3.0505
1.7275 | -34.0293 | -3.0661
1.7336 | -34.029 | -3.0821
1.7397 | -34.0291 | -3.0972
1.7458 | -34.0291 | -3.1137
1.7519 | -34.0289 | -3.1296
1.758 | -34.0288 | -3.1443
1.7641 | -34.0289 | -3.1587
1.7702 | -34.0287 | -3.1742
1.7763 | -34.0287 | -3.1877
1.7824 | -34.0285 | -3.2037
1.7885 | -34.0283 | -3.2176
1.7947 | -34.0284 | -3.2323
1.8008 | -34.0283 | -3.245
1.8069 | -34.0282 | -3.2585
1.813 | -34.0283 | -3.2724
1.8191 | -34.028 | -3.2865
1.8252 | -34.0277 | -3.3002
1.8313 | -34.0277 | -3.3144
1.8374 | -34.0277 | -3.3277
1.8435 | -34.0278 | -3.3396
1.8496 | -34.0275 | -3.3539
1.8557 | -34.0274 | -3.3684
1.8618 | -34.0273 | -3.38
1.8679 | -34.0273 | -3.3928
1.874 | -34.0273 | -3.4058
1.8801 | -34.0272 | -3.4185
1.8862 | -34.0273 | -3.4311
1.8923 | -34.0272 | -3.4431
1.8984 | -34.0269 | -3.4536
1.9045 | -34.0268 | -3.4645
1.9106 | -34.0265 | -3.4761
1.9167 | -34.0262 | -3.4871
1.9228 | -34.0261 | -3.4978
1.9289 | -34.0258 | -3.5065
1.9351 | -34.0258 | -3.5153
1.9412 | -34.0259 | -3.5255
1.9473 | -34.0259 | -3.5373
1.9534 | -34.0258 | -3.5482
1.9595 | -34.0256 | -3.5596
1.9656 | -34.0251 | -3.5701
1.9717 | -34.0248 | -3.5804
1.9778 | -34.0248 | -3.5918
1.9839 | -34.0248 | -3.603
1.99 | -34.0249 | -3.6127
1.9961 | -34.0246 | -3.6229
2.0022 | -34.0245 | -3.6327
2.0083 | -34.0241 | -3.6425
2.0144 | -34.0241 | -3.6522
2.0205 | -34.0241 | -3.6639
2.0266 | -34.0242 | -3.6753
2.0327 | -34.0242 | -3.6863
2.0388 | -34.0244 | -3.7
2.0449 | -34.0245 | -3.7102
2.051 | -34.0245 | -3.72
2.0571 | -34.0244 | -3.7312
2.0632 | -34.0242 | -3.7416
2.0693 | -34.0242 | -3.7526
2.0754 | -34.0244 | -3.7642
2.0816 | -34.0245 | -3.776
2.0877 | -34.0244 | -3.787
2.0938 | -34.0247 | -3.7991
2.0999 | -34.0246 | -3.8106
2.106 | -34.0248 | -3.8224
2.1121 | -34.0248 | -3.8344
2.1182 | -34.0249 | -3.8444
2.1243 | -34.0251 | -3.8548
2.1304 | -34.0252 | -3.8649
2.1365 | -34.0255 | -3.8763
2.1426 | -34.0256 | -3.8895
2.1487 | -34.0261 | -3.901
2.1548 | -34.0265 | -3.9139
2.1609 | -34.027 | -3.9258
2.167 | -34.0274 | -3.9383
2.1731 | -34.0277 | -3.95
2.1792 | -34.0284 | -3.9611
2.1853 | -34.0289 | -3.9738
2.1914 | -34.0294 | -3.9858
2.1975 | -34.03 | -3.9978
2.2036 | -34.0304 | -4.0109
2.2097 | -34.0308 | -4.0227
2.2158 | -34.0317 | -4.0329
2.222 | -34.0323 | -4.0429
2.2281 | -34.0326 | -4.0535
2.2342 | -34.0331 | -4.0665
2.2403 | -34.0335 | -4.0765
2.2464 | -34.0343 | -4.0875
2.2525 | -34.0347 | -4.0981
2.2586 | -34.0356 | -4.1079
2.2647 | -34.0364 | -4.1171
2.2708 | -34.037 | -4.1287
2.2769 | -34.0379 | -4.1401
2.283 | -34.0383 | -4.1493
2.2891 | -34.0389 | -4.1595
2.2952 | -34.0395 | -4.1692
2.3013 | -34.0401 | -4.1793
2.3074 | -34.0408 | -4.1888
2.3135 | -34.0413 | -4.2
2.3196 | -34.0419 | -4.2085
2.3257 | -34.0422 | -4.219
2.3318 | -34.0426 | -4.2278
2.3379 | -34.0431 | -4.2391
2.344 | -34.0434 | -4.2499
2.3501 | -34.0439 | -4.2605
2.3562 | -34.0442 | -4.2687
2.3623 | -34.0445 | -4.2773
2.3685 | -34.0447 | -4.2862
2.3746 | -34.045 | -4.2966
2.3807 | -34.0455 | -4.3054
2.3868 | -34.0456 | -4.3143
2.3929 | -34.046 | -4.3235
2.399 | -34.0463 | -4.3314
2.4051 | -34.0468 | -4.3408
2.4112 | -34.0471 | -4.3497
2.4173 | -34.0474 | -4.3585
2.4234 | -34.0474 | -4.3668
2.4295 | -34.0477 | -4.3753
2.4356 | -34.0478 | -4.3845
2.4417 | -34.0478 | -4.3944
2.4478 | -34.0479 | -4.4039
2.4539 | -34.0482 | -4.4116
2.46 | -34.0479 | -4.4212
2.4661 | -34.0479 | -4.4303
2.4722 | -34.0482 | -4.4389
2.4783 | -34.0483 | -4.4472
2.4844 | -34.0483 | -4.4556
2.4905 | -34.0482 | -4.4647
2.4966 | -34.0479 | -4.4735
2.5027 | -34.0479 | -4.4814
2.5089 | -34.0478 | -4.4898
2.515 | -34.0476 | -4.4989
2.5211 | -34.0475 | -4.5071
2.5272 | -34.0472 | -4.516
2.5333 | -34.047 | -4.5254
2.5394 | -34.047 | -4.5343
2.5455 | -34.0468 | -4.5436
2.5516 | -34.0467 | -4.5522
2.5577 | -34.0466 | -4.562
2.5638 | -34.0464 | -4.5722
2.5699 | -34.0462 | -4.5823
2.576 | -34.0459 | -4.5925
2.5821 | -34.0457 | -4.6023
2.5882 | -34.0455 | -4.6117
2.5943 | -34.0454 | -4.6214
2.6004 | -34.0452 | -4.6313
2.6065 | -34.045 | -4.6396
2.6126 | -34.0449 | -4.6486
2.6187 | -34.0448 | -4.6573
2.6248 | -34.0446 | -4.6662
2.6309 | -34.0445 | -4.6774
2.637 | -34.0445 | -4.6884
2.6431 | -34.0444 | -4.6978
2.6492 | -34.0443 | -4.7077
2.6554 | -34.0443 | -4.7178
2.6615 | -34.0441 | -4.7276
2.6676 | -34.044 | -4.7379
2.6737 | -34.0441 | -4.7485
2.6798 | -34.0441 | -4.7583
2.6859 | -34.0441 | -4.7687
2.692 | -34.044 | -4.7784
2.6981 | -34.0439 | -4.7894
2.7042 | -34.044 | -4.8013
2.7103 | -34.044 | -4.8113
2.7164 | -34.0442 | -4.8219
2.7225 | -34.0443 | -4.8331
2.7286 | -34.0443 | -4.8442
2.7347 | -34.0443 | -4.8541
2.7408 | -34.0446 | -4.8633
2.7469 | -34.0446 | -4.8746
2.753 | -34.0448 | -4.8851
2.7591 | -34.045 | -4.897
2.7652 | -34.045 | -4.9084
2.7713 | -34.045 | -4.9181
2.7774 | -34.0454 | -4.9294
2.7835 | -34.0455 | -4.94
2.7896 | -34.0456 | -4.9492
2.7958 | -34.046 | -4.9594
2.8019 | -34.0465 | -4.9692
2.808 | -34.0468 | -4.9809
2.8141 | -34.0472 | -4.9938
2.8202 | -34.0475 | -5.0059
2.8263 | -34.048 | -5.0167
2.8324 | -34.0482 | -5.029
2.8385 | -34.0486 | -5.0411
2.8446 | -34.0492 | -5.0527
2.8507 | -34.0495 | -5.0647
2.8568 | -34.05 | -5.0758
2.8629 | -34.0504 | -5.0871
2.869 | -34.0508 | -5.0983
2.8751 | -34.0512 | -5.1102
2.8812 | -34.0518 | -5.1234
2.8873 | -34.0522 | -5.1346
2.8934 | -34.0525 | -5.1464
2.8995 | -34.0528 | -5.1583
2.9056 | -34.0532 | -5.169
2.9117 | -34.0536 | -5.1805
2.9178 | -34.0543 | -5.1931
2.9239 | -34.0547 | -5.2063
2.93 | -34.0551 | -5.2197
2.9361 | -34.0557 | -5.231
2.9423 | -34.0563 | -5.2429
2.9484 | -34.0567 | -5.2564
2.9545 | -34.0574 | -5.2695
2.9606 | -34.0579 | -5.2828
2.9667 | -34.0584 | -5.2945
2.9728 | -34.0588 | -5.3067
2.9789 | -34.0592 | -5.3193
2.985 | -34.0597 | -5.333
2.9911 | -34.0601 | -5.3455
2.9972 | -34.0606 | -5.358
3.0033 | -34.0612 | -5.371
3.0094 | -34.0618 | -5.384
3.0155 | -34.0623 | -5.3965
3.0216 | -34.0628 | -5.4098
3.0277 | -34.0632 | -5.4234
3.0338 | -34.0638 | -5.4369
3.0399 | -34.0644 | -5.4506
3.046 | -34.0647 | -5.4635
3.0521 | -34.0653 | -5.4759
3.0582 | -34.0655 | -5.4896
3.0643 | -34.0661 | -5.5025
3.0704 | -34.0663 | -5.5158
3.0765 | -34.0667 | -5.5279
3.0827 | -34.067 | -5.5409
3.0888 | -34.0673 | -5.5537
3.0949 | -34.0676 | -5.5666
3.101 | -34.068 | -5.5806
3.1071 | -34.0683 | -5.5946
3.1132 | -34.0688 | -5.6085
3.1193 | -34.0692 | -5.6214
3.1254 | -34.0695 | -5.6335
3.1315 | -34.0699 | -5.6477
3.1376 | -34.0702 | -5.6605
3.1437 | -34.0705 | -5.6739
3.1498 | -34.0708 | -5.688
3.1559 | -34.0712 | -5.7011
3.162 | -34.0712 | -5.715
3.1681 | -34.0712 | -5.7291
3.1742 | -34.0716 | -5.7427
3.1803 | -34.0719 | -5.7558
3.1864 | -34.0723 | -5.768
3.1925 | -34.0726 | -5.7815
3.1986 | -34.0726 | -5.7942
3.2047 | -34.0727 | -5.808
3.2108 | -34.073 | -5.8196
3.2169 | -34.0734 | -5.8332
3.223 | -34.0733 | -5.8465
3.2292 | -34.0733 | -5.8603
3.2353 | -34.0732 | -5.8697
3.2414 | -34.0733 | -5.8826
3.2475 | -34.0733 | -5.8935
3.2536 | -34.0735 | -5.9071
3.2597 | -34.0737 | -5.9182
3.2658 | -34.0739 | -5.9288
3.2719 | -34.074 | -5.9398
3.278 | -34.074 | -5.9518
3.2841 | -34.0742 | -5.9636
3.2902 | -34.0742 | -5.976
3.2963 | -34.0744 | -5.9876
3.3024 | -34.0742 | -5.9976
3.3085 | -34.0745 | -6.01
3.3146 | -34.0743 | -6.0208
3.3207 | -34.0745 | -6.0328
3.3268 | -34.0746 | -6.0449
3.3329 | -34.0747 | -6.0546
3.339 | -34.0748 | -6.0641
3.3451 | -34.0748 | -6.0754
3.3512 | -34.0749 | -6.0847
3.3573 | -34.0751 | -6.0959
3.3634 | -34.0751 | -6.1073
3.3696 | -34.0751 | -6.1191
3.3757 | -34.0751 | -6.1288
3.3818 | -34.0754 | -6.1404
3.3879 | -34.0753 | -6.1506
3.394 | -34.0752 | -6.1609
3.4001 | -34.0752 | -6.1717
3.4062 | -34.0752 | -6.1817
3.4123 | -34.0753 | -6.1914
3.4184 | -34.0753 | -6.2023
3.4245 | -34.0754 | -6.2127
3.4306 | -34.0753 | -6.2223
3.4367 | -34.0754 | -6.2319
3.4428 | -34.0757 | -6.241
3.4489 | -34.0756 | -6.252
3.455 | -34.0755 | -6.2623
3.4611 | -34.0755 | -6.2733
3.4672 | -34.0755 | -6.2844
3.4733 | -34.0756 | -6.2944
3.4794 | -34.0755 | -6.3048
3.4855 | -34.0758 | -6.3143
3.4916 | -34.0759 | -6.3239
3.4977 | -34.0758 | -6.3347
3.5038 | -34.0759 | -6.3445
3.5099 | -34.076 | -6.354
3.5161 | -34.0761 | -6.3633
3.5222 | -34.0762 | -6.3732
3.5283 | -34.0761 | -6.3831
3.5344 | -34.0761 | -6.3938
3.5405 | -34.0763 | -6.4036
3.5466 | -34.0764 | -6.4132
3.5527 | -34.0764 | -6.4237
3.5588 | -34.0765 | -6.4333
3.5649 | -34.0765 | -6.4433
3.571 | -34.0766 | -6.4534
3.5771 | -34.0767 | -6.464
3.5832 | -34.0768 | -6.4735
3.5893 | -34.077 | -6.4832
3.5954 | -34.0771 | -6.4936
3.6015 | -34.0773 | -6.5038
3.6076 | -34.0775 | -6.5133
3.6137 | -34.0775 | -6.523
3.6198 | -34.0779 | -6.5319
3.6259 | -34.0779 | -6.5425
3.632 | -34.0782 | -6.553
3.6381 | -34.0783 | -6.5627
3.6442 | -34.0785 | -6.573
3.6503 | -34.0788 | -6.5819
3.6565 | -34.0791 | -6.5935
3.6626 | -34.0791 | -6.6041
3.6687 | -34.0794 | -6.6146
3.6748 | -34.0797 | -6.6243
3.6809 | -34.08 | -6.6347
3.687 | -34.0802 | -6.6457
3.6931 | -34.0806 | -6.6555
3.6992 | -34.0808 | -6.666
3.7053 | -34.081 | -6.676
3.7114 | -34.0812 | -6.6871
3.7175 | -34.0815 | -6.6969
3.7236 | -34.0817 | -6.7055
3.7297 | -34.0818 | -6.7161
3.7358 | -34.0821 | -6.7276
3.7419 | -34.0822 | -6.7391
3.748 | -34.0824 | -6.7497
3.7541 | -34.0825 | -6.7605
3.7602 | -34.0829 | -6.7712
3.7663 | -34.083 | -6.7812
3.7724 | -34.0834 | -6.7921
3.7785 | -34.0837 | -6.8024
3.7846 | -34.0839 | -6.8128
3.7907 | -34.0843 | -6.8237
3.7969 | -34.0845 | -6.835
3.803 | -34.0848 | -6.8449
3.8091 | -34.0851 | -6.8551
3.8152 | -34.0854 | -6.8662
3.8213 | -34.0856 | -6.8776
3.8274 | -34.0859 | -6.888
3.8335 | -34.0863 | -6.8994
3.8396 | -34.0868 | -6.9108
3.8457 | -34.0871 | -6.9209
3.8518 | -34.0876 | -6.9316
3.8579 | -34.0879 | -6.942
3.864 | -34.0883 | -6.9525
3.8701 | -34.0887 | -6.9633
3.8762 | -34.089 | -6.9744
3.8823 | -34.0892 | -6.9849
3.8884 | -34.0897 | -6.995
3.8945 | -34.0901 | -7.0062
3.9006 | -34.0905 | -7.0176
3.9067 | -34.0908 | -7.0293
3.9128 | -34.0912 | -7.0404
3.9189 | -34.0916 | -7.0509
3.925 | -34.0918 | -7.0616
3.9311 | -34.0922 | -7.0715
3.9372 | -34.0926 | -7.0829
3.9434 | -34.093 | -7.0932
3.9495 | -34.0933 | -7.1047
3.9556 | -34.0938 | -7.1157
3.9617 | -34.0941 | -7.1265
3.9678 | -34.0945 | -7.1371
3.9739 | -34.0949 | -7.1481
3.98 | -34.0952 | -7.1598
3.9861 | -34.0955 | -7.1701
3.9922 | -34.0957 | -7.1817
3.9983 | -34.0961 | -7.1916
4.0044 | -34.0964 | -7.2012
4.0105 | -34.0968 | -7.2119
4.0166 | -34.0973 | -7.2228
4.0227 | -34.0977 | -7.233
4.0288 | -34.0979 | -7.2442
4.0349 | -34.0982 | -7.254
4.041 | -34.0985 | -7.2648
4.0471 | -34.0988 | -7.2749
4.0532 | -34.0992 | -7.2859
4.0593 | -34.0995 | -7.2973
4.0654 | -34.0998 | -7.308
4.0715 | -34.1001 | -7.3192
4.0776 | -34.1005 | -7.3294
4.0838 | -34.1008 | -7.3406
4.0899 | -34.1011 | -7.3518
4.096 | -34.1014 | -7.362
4.1021 | -34.1017 | -7.3727
4.1082 | -34.1021 | -7.3829
4.1143 | -34.1024 | -7.3927
4.1204 | -34.1029 | -7.4033
4.1265 | -34.1032 | -7.4136
4.1326 | -34.1036 | -7.4238
4.1387 | -34.104 | -7.4349
4.1448 | -34.1043 | -7.4457
4.1509 | -34.1046 | -7.4563
4.157 | -34.1048 | -7.4674
4.1631 | -34.1051 | -7.4785
4.1692 | -34.1055 | -7.489
4.1753 | -34.1058 | -7.4998
4.1814 | -34.1061 | -7.5095
4.1875 | -34.1064 | -7.5207
4.1936 | -34.1066 | -7.5319
4.1997 | -34.107 | -7.542
4.2058 | -34.1072 | -7.5534
4.2119 | -34.1075 | -7.563
4.218 | -34.1077 | -7.573
4.2241 | -34.1081 | -7.5825
4.2303 | -34.1085 | -7.5936
4.2364 | -34.1088 | -7.6049
4.2425 | -34.1091 | -7.6153
4.2486 | -34.1093 | -7.6263
4.2547 | -34.1095 | -7.6355
4.2608 | -34.1097 | -7.6453
4.2669 | -34.11 | -7.6555
4.273 | -34.1103 | -7.6662
4.2791 | -34.1105 | -7.6768
4.2852 | -34.1107 | -7.6869
4.2913 | -34.1108 | -7.6961
4.2974 | -34.1111 | -7.7069
4.3035 | -34.1113 | -7.7175
4.3096 | -34.1114 | -7.7279
4.3157 | -34.1117 | -7.7377
4.3218 | -34.1119 | -7.7476
4.3279 | -34.1121 | -7.7578
4.334 | -34.1123 | -7.7676
4.3401 | -34.1125 | -7.7774
4.3462 | -34.1128 | -7.788
4.3523 | -34.1131 | -7.7977
4.3584 | -34.1133 | -7.806
4.3645 | -34.1134 | -7.8155
4.3707 | -34.1137 | -7.8247
4.3768 | -34.1139 | -7.835
4.3829 | -34.114 | -7.8452
4.389 | -34.1143 | -7.8548
4.3951 | -34.1145 | -7.8643
4.4012 | -34.1147 | -7.8739
4.4073 | -34.1149 | -7.8839
4.4134 | -34.1151 | -7.894
4.4195 | -34.1154 | -7.9035
4.4256 | -34.1156 | -7.913
4.4317 | -34.1158 | -7.9224
4.4378 | -34.116 | -7.9318
4.4439 | -34.1162 | -7.9412
4.45 | -34.1163 | -7.9511
4.4561 | -34.1164 | -7.9608
4.4622 | -34.1165 | -7.9703
4.4683 | -34.1167 | -7.9794
4.4744 | -34.1169 | -7.9895
4.4805 | -34.1171 | -7.9984
4.4866 | -34.1173 | -8.0076
4.4927 | -34.1175 | -8.0177
4.4988 | -34.1177 | -8.027
4.5049 | -34.1179 | -8.036
4.511 | -34.1181 | -8.0449
4.5172 | -34.1184 | -8.0546
4.5233 | -34.1186 | -8.0651
4.5294 | -34.1188 | -8.0743
4.5355 | -34.119 | -8.0836
4.5416 | -34.1193 | -8.0926
4.5477 | -34.1194 | -8.1018
4.5538 | -34.1196 | -8.1104
4.5599 | -34.1199 | -8.1198
4.566 | -34.1201 | -8.1294
4.5721 | -34.1203 | -8.1385
4.5782 | -34.1206 | -8.1482
4.5843 | -34.1208 | -8.1577
4.5904 | -34.1211 | -8.1669
4.5965 | -34.1214 | -8.1755
4.6026 | -34.1217 | -8.186
4.6087 | -34.1218 | -8.1949
4.6148 | -34.1221 | -8.2054
4.6209 | -34.1224 | -8.2138
4.627 | -34.1227 | -8.2237
4.6331 | -34.1228 | -8.2325
4.6392 | -34.123 | -8.2416
4.6453 | -34.1233 | -8.2509
4.6514 | -34.1236 | -8.2603
4.6576 | -34.1236 | -8.2707
4.6637 | -34.1238 | -8.2782
4.6698 | -34.124 | -8.2866
4.6759 | -34.1241 | -8.2963
4.682 | -34.1243 | -8.3053
4.6881 | -34.1244 | -8.314
4.6942 | -34.1247 | -8.3232
4.7003 | -34.1249 | -8.3325
4.7064 | -34.1251 | -8.3417
4.7125 | -34.1255 | -8.351
4.7186 | -34.1256 | -8.3593
4.7247 | -34.1259 | -8.3688
4.7308 | -34.1262 | -8.378
4.7369 | -34.1264 | -8.3875
4.743 | -34.1266 | -8.3977
4.7491 | -34.1267 | -8.4072
4.7552 | -34.1269 | -8.4163
4.7613 | -34.127 | -8.4258
4.7674 | -34.1273 | -8.4343
4.7735 | -34.1274 | -8.4442
4.7796 | -34.1276 | -8.4539
4.7857 | -34.1277 | -8.4635
4.7918 | -34.128 | -8.472
4.7979 | -34.1283 | -8.4809
4.8041 | -34.1284 | -8.4901
4.8102 | -34.1287 | -8.499
4.8163 | -34.1289 | -8.5084
4.8224 | -34.129 | -8.5169
4.8285 | -34.1294 | -8.5259
4.8346 | -34.1296 | -8.5353
4.8407 | -34.1298 | -8.5445
4.8468 | -34.13 | -8.5541
4.8529 | -34.1302 | -8.5638
4.859 | -34.1304 | -8.5731
4.8651 | -34.1307 | -8.5824
4.8712 | -34.1309 | -8.5916
4.8773 | -34.1311 | -8.6011
4.8834 | -34.1313 | -8.6107
4.8895 | -34.1316 | -8.6206
4.8956 | -34.1317 | -8.6305
4.9017 | -34.1319 | -8.6399
4.9078 | -34.1321 | -8.6496
4.9139 | -34.1323 | -8.6581
4.92 | -34.1326 | -8.667
4.9261 | -34.1329 | -8.6757
4.9322 | -34.1332 | -8.6848
4.9383 | -34.1336 | -8.6937
4.9445 | -34.1338 | -8.7037
4.9506 | -34.134 | -8.7124
4.9567 | -34.1343 | -8.7216
4.9628 | -34.1345 | -8.7309
4.9689 | -34.1347 | -8.7406
4.975 | -34.1349 | -8.7498
4.9811 | -34.1352 | -8.7595
4.9872 | -34.1354 | -8.7692
4.9933 | -34.1357 | -8.7787
4.9994 | -34.1361 | -8.7885
5.0055 | -34.1364 | -8.798
5.0116 | -34.1367 | -8.8081
5.0177 | -34.1371 | -8.8181
5.0238 | -34.1374 | -8.8284
5.0299 | -34.1378 | -8.8393
5.036 | -34.1381 | -8.8487
5.0421 | -34.1383 | -8.8583
5.0482 | -34.1385 | -8.8685
5.0543 | -34.1388 | -8.8777
5.0604 | -34.1392 | -8.8879
5.0665 | -34.1395 | -8.8974
5.0726 | -34.1398 | -8.9077
5.0787 | -34.1401 | -8.9178
5.0848 | -34.1404 | -8.9278
5.091 | -34.1407 | -8.9376
5.0971 | -34.141 | -8.9476
5.1032 | -34.1414 | -8.9578
5.1093 | -34.1417 | -8.9676
5.1154 | -34.1419 | -8.9775
5.1215 | -34.1421 | -8.9874
5.1276 | -34.1425 | -8.9971
5.1337 | -34.1429 | -9.0068
5.1398 | -34.1431 | -9.0164
5.1459 | -34.1435 | -9.0268
5.152 | -34.1438 | -9.0377
5.1581 | -34.144 | -9.0483
5.1642 | -34.1443 | -9.0581
5.1703 | -34.1446 | -9.068
5.1764 | -34.1449 | -9.0778
5.1825 | -34.1452 | -9.0881
5.1886 | -34.1455 | -9.0978
5.1947 | -34.1458 | -9.1085
5.2008 | -34.1462 | -9.1182
5.2069 | -34.1465 | -9.1292
5.213 | -34.1468 | -9.1387
5.2191 | -34.1472 | -9.1489
5.2252 | -34.1475 | -9.1591
5.2314 | -34.1477 | -9.1697
5.2375 | -34.1481 | -9.1796
5.2436 | -34.1483 | -9.1897
5.2497 | -34.1487 | -9.1994
5.2558 | -34.1491 | -9.2103
5.2619 | -34.1493 | -9.2211
5.268 | -34.1497 | -9.2314
5.2741 | -34.15 | -9.2425
5.2802 | -34.1503 | -9.2532
5.2863 | -34.1507 | -9.2633
5.2924 | -34.151 | -9.2744
5.2985 | -34.1514 | -9.2856
5.3046 | -34.1518 | -9.2967
5.3107 | -34.1522 | -9.3075
5.3168 | -34.1526 | -9.3178
5.3229 | -34.1529 | -9.3275
5.329 | -34.1533 | -9.3374
5.3351 | -34.1537 | -9.3475
5.3412 | -34.1541 | -9.3575
5.3473 | -34.1545 | -9.3683
5.3534 | -34.1549 | -9.3791
5.3595 | -34.1552 | -9.3897
5.3656 | -34.1555 | -9.4006
5.3717 | -34.1558 | -9.411
5.3779 | -34.1561 | -9.4218
5.384 | -34.1564 | -9.4324
5.3901 | -34.1568 | -9.443
5.3962 | -34.1571 | -9.4537
5.4023 | -34.1575 | -9.4644
5.4084 | -34.1579 | -9.4747
5.4145 | -34.1582 | -9.4856
5.4206 | -34.1584 | -9.4968
5.4267 | -34.1588 | -9.5081
5.4328 | -34.1592 | -9.5181
5.4389 | -34.1595 | -9.5282
5.445 | -34.1598 | -9.5395
5.4511 | -34.1602 | -9.5498
5.4572 | -34.1606 | -9.5608
5.4633 | -34.1609 | -9.5711
5.4694 | -34.1611 | -9.5815
5.4755 | -34.1615 | -9.5923
5.4816 | -34.1618 | -9.6028
5.4877 | -34.1622 | -9.6134
5.4938 | -34.1625 | -9.6238
5.4999 | -34.1628 | -9.6344
5.506 | -34.1632 | -9.645
5.5121 | -34.1635 | -9.656
5.5183 | -34.1638 | -9.666
5.5244 | -34.1641 | -9.6767
5.5305 | -34.1644 | -9.6875
5.5366 | -34.1648 | -9.6978
5.5427 | -34.1651 | -9.709
5.5488 | -34.1654 | -9.7199
5.5549 | -34.1656 | -9.7305
5.561 | -34.166 | -9.7417
5.5671 | -34.1664 | -9.752
5.5732 | -34.1666 | -9.7628
5.5793 | -34.167 | -9.7743
5.5854 | -34.1674 | -9.7851
5.5915 | -34.1676 | -9.7959
5.5976 | -34.1679 | -9.8068
5.6037 | -34.1683 | -9.8174
5.6098 | -34.1685 | -9.8279
5.6159 | -34.1689 | -9.8387
5.622 | -34.1693 | -9.8491
5.6281 | -34.1696 | -9.8598
5.6342 | -34.1699 | -9.8706
5.6403 | -34.1702 | -9.8807
5.6464 | -34.1706 | -9.8914
5.6525 | -34.1708 | -9.9018
5.6586 | -34.1712 | -9.9117
5.6648 | -34.1715 | -9.9229
5.6709 | -34.1718 | -9.9339
5.677 | -34.1722 | -9.9453
5.6831 | -34.1725 | -9.9554
5.6892 | -34.1728 | -9.9661
5.6953 | -34.1731 | -9.9767
5.7014 | -34.1734 | -9.9879
5.7075 | -34.1737 | -9.9987
5.7136 | -34.174 | -10.0091
5.7197 | -34.1743 | -10.0197
5.7258 | -34.1746 | -10.0301
5.7319 | -34.1749 | -10.0408
5.738 | -34.1753 | -10.0514
5.7441 | -34.1756 | -10.062
5.7502 | -34.176 | -10.0722
5.7563 | -34.1763 | -10.0826
5.7624 | -34.1766 | -10.094
5.7685 | -34.1768 | -10.1044
5.7746 | -34.177 | -10.1148
5.7807 | -34.1774 | -10.1251
5.7868 | -34.1776 | -10.1356
5.7929 | -34.1779 | -10.1461
5.799 | -34.1782 | -10.1566
5.8052 | -34.1785 | -10.1673
5.8113 | -34.1788 | -10.1779
5.8174 | -34.179 | -10.1886
5.8235 | -34.1793 | -10.199
5.8296 | -34.1796 | -10.2101
5.8357 | -34.1798 | -10.2207
5.8418 | -34.1802 | -10.2314
5.8479 | -34.1804 | -10.2423
5.854 | -34.1806 | -10.2531
5.8601 | -34.181 | -10.2635
5.8662 | -34.1812 | -10.2744
5.8723 | -34.1815 | -10.2844
5.8784 | -34.1818 | -10.2949
5.8845 | -34.1821 | -10.3056
5.8906 | -34.1823 | -10.316
5.8967 | -34.1827 | -10.3262
5.9028 | -34.183 | -10.3368
5.9089 | -34.1832 | -10.3467
5.915 | -34.1835 | -10.3572
5.9211 | -34.1839 | -10.3678
5.9272 | -34.1841 | -10.3782
5.9333 | -34.1844 | -10.3881
5.9394 | -34.1847 | -10.3984
5.9455 | -34.185 | -10.4085
5.9517 | -34.1854 | -10.419
5.9578 | -34.1857 | -10.4297
5.9639 | -34.186 | -10.4399
5.97 | -34.1864 | -10.4505
5.9761 | -34.1867 | -10.4606
5.9822 | -34.187 | -10.4712
5.9883 | -34.1873 | -10.4814
5.9944 | -34.1876 | -10.4919
6.0005 | -34.188 | -10.5026
6.0066 | -34.1882 | -10.5126
6.0127 | -34.1885 | -10.5224
6.0188 | -34.1888 | -10.5331
6.0249 | -34.1891 | -10.5435
6.031 | -34.1894 | -10.5536
6.0371 | -34.1898 | -10.5643
6.0432 | -34.1902 | -10.5747
6.0493 | -34.1905 | -10.5849
6.0554 | -34.1909 | -10.5952
6.0615 | -34.1912 | -10.6061
6.0676 | -34.1915 | -10.6171
6.0737 | -34.1918 | -10.6277
6.0798 | -34.192 | -10.6379
6.0859 | -34.1924 | -10.6482
6.0921 | -34.1927 | -10.6584
6.0982 | -34.1931 | -10.669
6.1043 | -34.1935 | -10.6792
6.1104 | -34.1938 | -10.6901
6.1165 | -34.1941 | -10.6997
6.1226 | -34.1945 | -10.7095
6.1287 | -34.1948 | -10.7199
6.1348 | -34.1952 | -10.7303
6.1409 | -34.1955 | -10.7405
6.147 | -34.1958 | -10.7507
6.1531 | -34.1962 | -10.761
6.1592 | -34.1965 | -10.7713
6.1653 | -34.1968 | -10.7814
6.1714 | -34.1971 | -10.7918
6.1775 | -34.1975 | -10.8024
6.1836 | -34.1977 | -10.8125
6.1897 | -34.1981 | -10.8228
6.1958 | -34.1983 | -10.8333
6.2019 | -34.1987 | -10.8439
6.208 | -34.1991 | -10.8546
6.2141 | -34.1994 | -10.8651
6.2202 | -34.1998 | -10.8759
6.2263 | -34.2001 | -10.8861
6.2325 | -34.2005 | -10.8959
6.2386 | -34.2009 | -10.9059
6.2447 | -34.2011 | -10.9164
6.2508 | -34.2015 | -10.9266
6.2569 | -34.2018 | -10.9372
6.263 | -34.2022 | -10.9472
6.2691 | -34.2026 | -10.9574
6.2752 | -34.2029 | -10.9675
6.2813 | -34.2032 | -10.9778
6.2874 | -34.2035 | -10.9881
6.2935 | -34.2039 | -10.9982
6.2996 | -34.2042 | -11.0085
6.3057 | -34.2046 | -11.0187
6.3118 | -34.2049 | -11.0289
6.3179 | -34.2053 | -11.0391
6.324 | -34.2056 | -11.0496
6.3301 | -34.206 | -11.0597
6.3362 | -34.2064 | -11.07
6.3423 | -34.2068 | -11.0803
6.3484 | -34.2072 | -11.0908
6.3545 | -34.2075 | -11.1005
6.3606 | -34.2079 | -11.1104
6.3667 | -34.2083 | -11.1204
6.3728 | -34.2087 | -11.1307
6.379 | -34.209 | -11.1409
6.3851 | -34.2094 | -11.1512
6.3912 | -34.2097 | -11.1612
6.3973 | -34.2101 | -11.1711
6.4034 | -34.2105 | -11.1813
6.4095 | -34.2108 | -11.1918
6.4156 | -34.2111 | -11.2017
6.4217 | -34.2115 | -11.2118
6.4278 | -34.2118 | -11.2219
6.4339 | -34.2122 | -11.2321
6.44 | -34.2126 | -11.2419
6.4461 | -34.2129 | -11.2526
6.4522 | -34.2133 | -11.2622
6.4583 | -34.2136 | -11.2724
6.4644 | -34.2139 | -11.283
6.4705 | -34.2143 | -11.2933
6.4766 | -34.2147 | -11.3039
6.4827 | -34.215 | -11.3145
6.4888 | -34.2153 | -11.3246
6.4949 | -34.2157 | -11.3343
6.501 | -34.216 | -11.3448
6.5071 | -34.2164 | -11.3552
6.5132 | -34.2168 | -11.3652
6.5194 | -34.2172 | -11.3753
6.5255 | -34.2175 | -11.3851
6.5316 | -34.2178 | -11.3954
6.5377 | -34.2182 | -11.4059
6.5438 | -34.2186 | -11.4159
6.5499 | -34.219 | -11.4259
6.556 | -34.2193 | -11.4359
6.5621 | -34.2196 | -11.4463
6.5682 | -34.2199 | -11.4562
6.5743 | -34.2203 | -11.4672
6.5804 | -34.2206 | -11.4774
6.5865 | -34.2209 | -11.4878
6.5926 | -34.2213 | -11.4979
6.5987 | -34.2217 | -11.5077
6.6048 | -34.222 | -11.5186
6.6109 | -34.2224 | -11.5287
6.617 | -34.2227 | -11.5397
6.6231 | -34.223 | -11.5501
6.6292 | -34.2234 | -11.5602
6.6353 | -34.2238 | -11.5714
6.6414 | -34.2242 | -11.5819
6.6475 | -34.2246 | -11.5923
6.6536 | -34.2249 | -11.6021
6.6597 | -34.2253 | -11.6121
6.6659 | -34.2257 | -11.6221
6.672 | -34.2261 | -11.6324
6.6781 | -34.2264 | -11.6424
6.6842 | -34.2268 | -11.6527
6.6903 | -34.2272 | -11.6628
6.6964 | -34.2275 | -11.6731
6.7025 | -34.2278 | -11.6833
6.7086 | -34.2282 | -11.6933
6.7147 | -34.2285 | -11.7036
6.7208 | -34.2289 | -11.7143
6.7269 | -34.2292 | -11.7241
6.733 | -34.2295 | -11.7335
6.7391 | -34.2299 | -11.7439
6.7452 | -34.2304 | -11.7536
6.7513 | -34.2307 | -11.7634
6.7574 | -34.2311 | -11.7738
6.7635 | -34.2315 | -11.7835
6.7696 | -34.2319 | -11.7938
6.7757 | -34.2323 | -11.8042
6.7818 | -34.2327 | -11.8145
6.7879 | -34.2331 | -11.8247
6.794 | -34.2334 | -11.8347
6.8001 | -34.2337 | -11.8446
6.8063 | -34.2341 | -11.855
6.8124 | -34.2345 | -11.8652
6.8185 | -34.2348 | -11.8753
6.8246 | -34.2352 | -11.8855
6.8307 | -34.2355 | -11.8955
6.8368 | -34.2359 | -11.906
6.8429 | -34.2363 | -11.916
6.849 | -34.2366 | -11.9262
6.8551 | -34.237 | -11.9364
6.8612 | -34.2374 | -11.9463
6.8673 | -34.2377 | -11.9563
6.8734 | -34.2381 | -11.9672
6.8795 | -34.2385 | -11.9775
6.8856 | -34.2389 | -11.9872
6.8917 | -34.2392 | -11.9967
6.8978 | -34.2396 | -12.0068
6.9039 | -34.2399 | -12.0167
6.91 | -34.2403 | -12.027
6.9161 | -34.2407 | -12.0369
6.9222 | -34.241 | -12.047
6.9283 | -34.2413 | -12.0568
6.9344 | -34.2417 | -12.0667
6.9405 | -34.242 | -12.0766
6.9466 | -34.2424 | -12.0869
6.9528 | -34.2427 | -12.097
6.9589 | -34.2431 | -12.107
6.965 | -34.2435 | -12.1171
6.9711 | -34.2439 | -12.1271
6.9772 | -34.2443 | -12.137
6.9833 | -34.2445 | -12.1472
6.9894 | -34.2449 | -12.1577
6.9955 | -34.2453 | -12.1678
7.0016 | -34.2456 | -12.1782
7.0077 | -34.246 | -12.1886
7.0138 | -34.2464 | -12.199
7.0199 | -34.2467 | -12.2089
7.026 | -34.247 | -12.2188
7.0321 | -34.2473 | -12.2289
7.0382 | -34.2476 | -12.2387
7.0443 | -34.248 | -12.2491
7.0504 | -34.2483 | -12.2595
7.0565 | -34.2487 | -12.2698
7.0626 | -34.249 | -12.2797
7.0687 | -34.2494 | -12.2894
7.0748 | -34.2497 | -12.2993
7.0809 | -34.2501 | -12.3095
7.087 | -34.2504 | -12.3196
7.0932 | -34.2507 | -12.3296
7.0993 | -34.251 | -12.3397
7.1054 | -34.2514 | -12.3495
7.1115 | -34.2517 | -12.3594
7.1176 | -34.252 | -12.3693
7.1237 | -34.2522 | -12.3793
7.1298 | -34.2525 | -12.3889
7.1359 | -34.2528 | -12.3986
7.142 | -34.2531 | -12.4087
7.1481 | -34.2535 | -12.4184
7.1542 | -34.2538 | -12.4283
7.1603 | -34.2541 | -12.4382
7.1664 | -34.2545 | -12.4482
7.1725 | -34.2548 | -12.458
7.1786 | -34.2552 | -12.4681
7.1847 | -34.2555 | -12.4779
7.1908 | -34.2558 | -12.488
7.1969 | -34.2561 | -12.4979
7.203 | -34.2565 | -12.5079
7.2091 | -34.2568 | -12.5176
7.2152 | -34.2571 | -12.5276
7.2213 | -34.2574 | -12.5373
7.2274 | -34.2578 | -12.5472
7.2335 | -34.2581 | -12.5571
7.2397 | -34.2584 | -12.567
7.2458 | -34.2587 | -12.5765
7.2519 | -34.259 | -12.5864
7.258 | -34.2594 | -12.5965
7.2641 | -34.2597 | -12.6063
7.2702 | -34.26 | -12.6165
7.2763 | -34.2604 | -12.6261
7.2824 | -34.2608 | -12.636
7.2885 | -34.2611 | -12.6458
7.2946 | -34.2615 | -12.6559
7.3007 | -34.2617 | -12.6658
7.3068 | -34.262 | -12.6759
7.3129 | -34.2623 | -12.6856
7.319 | -34.2627 | -12.6956
7.3251 | -34.263 | -12.7051
7.3312 | -34.2634 | -12.7146
7.3373 | -34.2637 | -12.7242
7.3434 | -34.2641 | -12.7338
7.3495 | -34.2645 | -12.7432
7.3556 | -34.2648 | -12.7535
7.3617 | -34.2651 | -12.7637
7.3678 | -34.2654 | -12.7737
7.3739 | -34.2658 | -12.7831
7.3801 | -34.266 | -12.7928
7.3862 | -34.2663 | -12.8026
7.3923 | -34.2667 | -12.8124
7.3984 | -34.267 | -12.8221
7.4045 | -34.2673 | -12.8318
7.4106 | -34.2677 | -12.8414
7.4167 | -34.268 | -12.8514
7.4228 | -34.2683 | -12.8612
7.4289 | -34.2686 | -12.8706
7.435 | -34.269 | -12.8802
7.4411 | -34.2694 | -12.8897
7.4472 | -34.2697 | -12.8995
7.4533 | -34.2701 | -12.9093
7.4594 | -34.2704 | -12.9194
7.4655 | -34.2707 | -12.9291
7.4716 | -34.271 | -12.9391
7.4777 | -34.2713 | -12.9484
7.4838 | -34.2717 | -12.9583
7.4899 | -34.2721 | -12.9681
7.496 | -34.2724 | -12.9779
7.5021 | -34.2727 | -12.9875
7.5082 | -34.273 | -12.9974
7.5143 | -34.2734 | -13.007
7.5204 | -34.2737 | -13.0164
7.5266 | -34.274 | -13.0262
7.5327 | -34.2743 | -13.0358
7.5388 | -34.2746 | -13.0455
7.5449 | -34.2749 | -13.0554
7.551 | -34.2752 | -13.0654
7.5571 | -34.2756 | -13.075
7.5632 | -34.2759 | -13.0847
7.5693 | -34.2762 | -13.0945
7.5754 | -34.2766 | -13.1043
7.5815 | -34.2769 | -13.1136
7.5876 | -34.2772 | -13.1233
7.5937 | -34.2775 | -13.1329
7.5998 | -34.2778 | -13.1427
7.6059 | -34.2781 | -13.1524
7.612 | -34.2785 | -13.162
7.6181 | -34.2788 | -13.1716
7.6242 | -34.2791 | -13.1814
7.6303 | -34.2795 | -13.1909
7.6364 | -34.2798 | -13.2008
7.6425 | -34.2801 | -13.2108
7.6486 | -34.2804 | -13.2206
7.6547 | -34.2808 | -13.2304
7.6608 | -34.2812 | -13.2401
7.667 | -34.2814 | -13.2492
7.6731 | -34.2817 | -13.2591
7.6792 | -34.282 | -13.2689
7.6853 | -34.2823 | -13.2785
7.6914 | -34.2826 | -13.2883
7.6975 | -34.2829 | -13.2982
7.7036 | -34.2832 | -13.308
7.7097 | -34.2835 | -13.3177
7.7158 | -34.2838 | -13.3272
7.7219 | -34.2841 | -13.3368
7.728 | -34.2844 | -13.3466
7.7341 | -34.2847 | -13.3564
7.7402 | -34.2851 | -13.366
7.7463 | -34.2853 | -13.3757
7.7524 | -34.2856 | -13.3857
7.7585 | -34.2859 | -13.3956
7.7646 | -34.2862 | -13.405
7.7707 | -34.2865 | -13.4145
7.7768 | -34.2867 | -13.4241
7.7829 | -34.287 | -13.4339
7.789 | -34.2873 | -13.4433
7.7951 | -34.2876 | -13.4533
7.8012 | -34.2879 | -13.4634
7.8073 | -34.2882 | -13.4729
7.8135 | -34.2884 | -13.4827
7.8196 | -34.2888 | -13.4924
7.8257 | -34.2891 | -13.5024
7.8318 | -34.2894 | -13.5125
7.8379 | -34.2897 | -13.5218
7.844 | -34.2899 | -13.531
7.8501 | -34.2902 | -13.541
7.8562 | -34.2904 | -13.5513
7.8623 | -34.2908 | -13.561
7.8684 | -34.291 | -13.5704
7.8745 | -34.2913 | -13.58
7.8806 | -34.2916 | -13.5892
7.8867 | -34.2919 | -13.5987
7.8928 | -34.2923 | -13.6083
7.8989 | -34.2926 | -13.6175
7.905 | -34.2929 | -13.6276
7.9111 | -34.2932 | -13.6373
7.9172 | -34.2935 | -13.6467
7.9233 | -34.2937 | -13.6564
7.9294 | -34.2941 | -13.6663
7.9355 | -34.2944 | -13.6755
7.9416 | -34.2946 | -13.6851
7.9477 | -34.2949 | -13.6943
7.9539 | -34.2952 | -13.7036
7.96 | -34.2955 | -13.7135
7.9661 | -34.2958 | -13.7233
7.9722 | -34.2962 | -13.733
7.9783 | -34.2964 | -13.7427
7.9844 | -34.2968 | -13.7528
7.9905 | -34.297 | -13.7623
7.9966 | -34.2973 | -13.7719
8.0027 | -34.2976 | -13.7819
8.0088 | -34.2979 | -13.7912
8.0149 | -34.2981 | -13.8006
8.021 | -34.2984 | -13.8103
8.0271 | -34.2987 | -13.8204
8.0332 | -34.299 | -13.83
8.0393 | -34.2993 | -13.8399
8.0454 | -34.2996 | -13.8497
8.0515 | -34.2999 | -13.8593
8.0576 | -34.3001 | -13.8693
8.0637 | -34.3004 | -13.8784
8.0698 | -34.3007 | -13.8876
8.0759 | -34.301 | -13.897
8.082 | -34.3013 | -13.9068
8.0881 | -34.3016 | -13.9163
8.0942 | -34.3018 | -13.9264
8.1004 | -34.3021 | -13.9364
8.1065 | -34.3024 | -13.946
8.1126 | -34.3028 | -13.9556
8.1187 | -34.3031 | -13.9653
8.1248 | -34.3034 | -13.9752
8.1309 | -34.3036 | -13.9848
8.137 | -34.3039 | -13.9944
8.1431 | -34.3042 | -14.0042
8.1492 | -34.3045 | -14.014
8.1553 | -34.3048 | -14.0237
8.1614 | -34.3051 | -14.0336
8.1675 | -34.3054 | -14.0431
8.1736 | -34.3057 | -14.0531
8.1797 | -34.3059 | -14.0629
8.1858 | -34.3062 | -14.0725
8.1919 | -34.3064 | -14.0821
8.198 | -34.3067 | -14.0914
8.2041 | -34.3071 | -14.1014
8.2102 | -34.3074 | -14.1111
8.2163 | -34.3077 | -14.1207
8.2224 | -34.308 | -14.1306
8.2285 | -34.3083 | -14.14
8.2346 | -34.3086 | -14.1496
8.2408 | -34.3089 | -14.1594
8.2469 | -34.3092 | -14.1689
8.253 | -34.3094 | -14.1785
8.2591 | -34.3097 | -14.1879
8.2652 | -34.3101 | -14.1974
8.2713 | -34.3104 | -14.2071
8.2774 | -34.3107 | -14.2169
8.2835 | -34.311 | -14.2269
8.2896 | -34.3113 | -14.2361
8.2957 | -34.3116 | -14.2457
8.3018 | -34.3119 | -14.2554
8.3079 | -34.3122 | -14.265
8.314 | -34.3125 | -14.2745
8.3201 | -34.3128 | -14.284
8.3262 | -34.3131 | -14.2936
8.3323 | -34.3134 | -14.3032
8.3384 | -34.3137 | -14.3133
8.3445 | -34.314 | -14.3225
8.3506 | -34.3143 | -14.3324
8.3567 | -34.3146 | -14.3422
8.3628 | -34.3149 | -14.3518
8.3689 | -34.3152 | -14.3615
8.375 | -34.3154 | -14.3712
8.3812 | -34.3158 | -14.3808
8.3873 | -34.3161 | -14.3905
8.3934 | -34.3164 | -14.4001
8.3995 | -34.3167 | -14.4099
8.4056 | -34.317 | -14.4196
8.4117 | -34.3172 | -14.4295
8.4178 | -34.3176 | -14.4388
8.4239 | -34.3178 | -14.4482
8.43 | -34.3182 | -14.4578
8.4361 | -34.3185 | -14.4679
8.4422 | -34.3188 | -14.4775
8.4483 | -34.319 | -14.4871
8.4544 | -34.3193 | -14.497
8.4605 | -34.3196 | -14.5066
8.4666 | -34.3199 | -14.5163
8.4727 | -34.3201 | -14.5259
8.4788 | -34.3205 | -14.5354
8.4849 | -34.3207 | -14.5454
8.491 | -34.321 | -14.5549
8.4971 | -34.3213 | -14.5647
8.5032 | -34.3216 | -14.5743
8.5093 | -34.3219 | -14.584
8.5154 | -34.3222 | -14.5934
8.5215 | -34.3225 | -14.6031
8.5277 | -34.3228 | -14.6129
8.5338 | -34.3231 | -14.6223
8.5399 | -34.3234 | -14.6322
8.546 | -34.3237 | -14.6416
8.5521 | -34.324 | -14.6513
8.5582 | -34.3243 | -14.6611
8.5643 | -34.3246 | -14.6706
8.5704 | -34.3249 | -14.6799
8.5765 | -34.3252 | -14.6895
8.5826 | -34.3256 | -14.6991
8.5887 | -34.3258 | -14.7086
8.5948 | -34.3261 | -14.7179
8.6009 | -34.3264 | -14.7275
8.607 | -34.3267 | -14.7373
8.6131 | -34.327 | -14.7472
8.6192 | -34.3273 | -14.7565
8.6253 | -34.3276 | -14.7663
8.6314 | -34.3279 | -14.7759
8.6375 | -34.3282 | -14.7857
8.6436 | -34.3286 | -14.7954
8.6497 | -34.3289 | -14.8051
8.6558 | -34.3291 | -14.8145
8.6619 | -34.3294 | -14.8243
8.6681 | -34.3297 | -14.8338
8.6742 | -34.33 | -14.8437
8.6803 | -34.3303 | -14.8536
8.6864 | -34.3307 | -14.8635
8.6925 | -34.331 | -14.8732
8.6986 | -34.3313 | -14.8831
8.7047 | -34.3316 | -14.8928
8.7108 | -34.3319 | -14.9026
8.7169 | -34.3322 | -14.912
8.723 | -34.3326 | -14.9216
8.7291 | -34.3328 | -14.9312
8.7352 | -34.3332 | -14.9411
8.7413 | -34.3335 | -14.9507
8.7474 | -34.3338 | -14.96
8.7535 | -34.3342 | -14.9698
8.7596 | -34.3345 | -14.98
8.7657 | -34.3348 | -14.9896
8.7718 | -34.3351 | -14.999
8.7779 | -34.3354 | -15.0092
8.784 | -34.3356 | -15.0187
8.7901 | -34.3359 | -15.0288
8.7962 | -34.3363 | -15.0385
8.8023 | -34.3366 | -15.0479
8.8084 | -34.3369 | -15.0577
8.8146 | -34.3372 | -15.0672
8.8207 | -34.3376 | -15.0769
8.8268 | -34.3379 | -15.0865
8.8329 | -34.3382 | -15.0962
8.839 | -34.3386 | -15.106
8.8451 | -34.3389 | -15.1155
8.8512 | -34.3392 | -15.1252
8.8573 | -34.3395 | -15.1346
8.8634 | -34.3399 | -15.1442
8.8695 | -34.3402 | -15.154
8.8756 | -34.3405 | -15.1634
8.8817 | -34.3408 | -15.1733
8.8878 | -34.3411 | -15.1832
8.8939 | -34.3414 | -15.1926
8.9 | -34.3417 | -15.2019
8.9061 | -34.342 | -15.2116
8.9122 | -34.3424 | -15.2211
8.9183 | -34.3427 | -15.2307
8.9244 | -34.343 | -15.2401
8.9305 | -34.3434 | -15.2498
8.9366 | -34.3437 | -15.2593
8.9427 | -34.344 | -15.2691
8.9488 | -34.3444 | -15.2787
8.955 | -34.3447 | -15.2883
8.9611 | -34.345 | -15.2981
8.9672 | -34.3454 | -15.3081
8.9733 | -34.3457 | -15.3176
8.9794 | -34.346 | -15.3274
8.9855 | -34.3464 | -15.337
8.9916 | -34.3467 | -15.3465
8.9977 | -34.347 | -15.356
9.0038 | -34.3474 | -15.3659
9.0099 | -34.3477 | -15.3756
9.016 | -34.3481 | -15.3853
9.0221 | -34.3484 | -15.3952
9.0282 | -34.3488 | -15.4049
9.0343 | -34.3491 | -15.4149
9.0404 | -34.3495 | -15.4244
9.0465 | -34.3498 | -15.4344
9.0526 | -34.3502 | -15.4439
9.0587 | -34.3505 | -15.4535
9.0648 | -34.3509 | -15.4631
9.0709 | -34.3513 | -15.4725
9.077 | -34.3516 | -15.4828
9.0831 | -34.3519 | -15.4924
9.0892 | -34.3522 | -15.5021
9.0953 | -34.3525 | -15.5119
9.1015 | -34.3528 | -15.5216
9.1076 | -34.3532 | -15.5313
9.1137 | -34.3536 | -15.5409
9.1198 | -34.3539 | -15.5502
9.1259 | -34.3543 | -15.5599
9.132 | -34.3547 | -15.5698
9.1381 | -34.3551 | -15.5794
9.1442 | -34.3555 | -15.5886
9.1503 | -34.3558 | -15.5984
9.1564 | -34.3561 | -15.6081
9.1625 | -34.3564 | -15.6178
9.1686 | -34.3568 | -15.6275
9.1747 | -34.3571 | -15.6373
9.1808 | -34.3575 | -15.6472
9.1869 | -34.3578 | -15.6571
9.193 | -34.3582 | -15.6669
9.1991 | -34.3585 | -15.6763
9.2052 | -34.3589 | -15.6859
9.2113 | -34.3593 | -15.6958
9.2174 | -34.3596 | -15.7061
9.2235 | -34.36 | -15.7155
9.2296 | -34.3604 | -15.7253
9.2357 | -34.3607 | -15.735
9.2419 | -34.3611 | -15.7447
9.248 | -34.3615 | -15.7546
9.2541 | -34.3618 | -15.7644
9.2602 | -34.3621 | -15.7742
9.2663 | -34.3625 | -15.7837
9.2724 | -34.3629 | -15.7934
9.2785 | -34.3632 | -15.8031
9.2846 | -34.3636 | -15.8129
9.2907 | -34.3639 | -15.8225
9.2968 | -34.3642 | -15.8323
9.3029 | -34.3646 | -15.8421
9.309 | -34.3649 | -15.8518
9.3151 | -34.3653 | -15.8613
9.3212 | -34.3657 | -15.8712
9.3273 | -34.3661 | -15.8808
9.3334 | -34.3665 | -15.8902
9.3395 | -34.3669 | -15.8999
9.3456 | -34.3673 | -15.9093
9.3517 | -34.3677 | -15.9186
9.3578 | -34.3681 | -15.9281
9.3639 | -34.3685 | -15.938
9.37 | -34.3688 | -15.9478
9.3761 | -34.3692 | -15.9573
9.3822 | -34.3695 | -15.9669
9.3884 | -34.3699 | -15.9765
9.3945 | -34.3703 | -15.9863
9.4006 | -34.3707 | -15.996
9.4067 | -34.371 | -16.0058
9.4128 | -34.3714 | -16.0154
9.4189 | -34.3718 | -16.0252
9.425 | -34.3722 | -16.0348
9.4311 | -34.3725 | -16.0444
9.4372 | -34.3729 | -16.0541
9.4433 | -34.3733 | -16.0639
9.4494 | -34.3736 | -16.0735
9.4555 | -34.374 | -16.0832
9.4616 | -34.3744 | -16.0928
9.4677 | -34.3748 | -16.1023
9.4738 | -34.3752 | -16.1124
9.4799 | -34.3755 | -16.1219
9.486 | -34.3758 | -16.1319
9.4921 | -34.3763 | -16.142
9.4982 | -34.3767 | -16.1513
9.5043 | -34.377 | -16.1607
9.5104 | -34.3774 | -16.1702
9.5165 | -34.3778 | -16.1799
9.5226 | -34.3782 | -16.1897
9.5288 | -34.3786 | -16.1993
9.5349 | -34.379 | -16.209
9.541 | -34.3793 | -16.2186
9.5471 | -34.3797 | -16.2283
9.5532 | -34.3801 | -16.238
9.5593 | -34.3805 | -16.2476
9.5654 | -34.3809 | -16.2573
9.5715 | -34.3813 | -16.2672
9.5776 | -34.3817 | -16.2772
9.5837 | -34.382 | -16.2865
9.5898 | -34.3824 | -16.2964
9.5959 | -34.3827 | -16.3059
9.602 | -34.3831 | -16.3157
9.6081 | -34.3834 | -16.3255
9.6142 | -34.3839 | -16.3351
9.6203 | -34.3842 | -16.3448
9.6264 | -34.3845 | -16.3541
9.6325 | -34.3848 | -16.364
9.6386 | -34.3853 | -16.3736
9.6447 | -34.3856 | -16.3832
9.6508 | -34.386 | -16.393
9.6569 | -34.3864 | -16.4026
9.663 | -34.3868 | -16.4123
9.6691 | -34.3872 | -16.422
9.6753 | -34.3876 | -16.4316
9.6814 | -34.388 | -16.441
9.6875 | -34.3885 | -16.4505
9.6936 | -34.3889 | -16.46
9.6997 | -34.3892 | -16.4698
9.7058 | -34.3896 | -16.4794
9.7119 | -34.39 | -16.489
9.718 | -34.3904 | -16.4989
9.7241 | -34.3907 | -16.5086
9.7302 | -34.3911 | -16.5183
9.7363 | -34.3915 | -16.5278
9.7424 | -34.3919 | -16.5374
9.7485 | -34.3922 | -16.547
9.7546 | -34.3927 | -16.5564
9.7607 | -34.393 | -16.5662
9.7668 | -34.3934 | -16.5763
9.7729 | -34.3938 | -16.5857
9.779 | -34.3943 | -16.5953
9.7851 | -34.3946 | -16.6051
9.7912 | -34.395 | -16.6153
9.7973 | -34.3953 | -16.6252
9.8034 | -34.3957 | -16.6351
9.8095 | -34.3961 | -16.6446
9.8157 | -34.3965 | -16.6543
9.8218 | -34.3969 | -16.6636
9.8279 | -34.3973 | -16.6733
9.834 | -34.3977 | -16.6829
9.8401 | -34.3981 | -16.6925
9.8462 | -34.3985 | -16.7018
9.8523 | -34.3989 | -16.7114
9.8584 | -34.3993 | -16.7212
9.8645 | -34.3997 | -16.7308
9.8706 | -34.4 | -16.7404
9.8767 | -34.4004 | -16.7496
9.8828 | -34.4008 | -16.7589
9.8889 | -34.4012 | -16.7688
9.895 | -34.4016 | -16.7781
9.9011 | -34.402 | -16.7872
9.9072 | -34.4024 | -16.7965
9.9133 | -34.4028 | -16.8059
9.9194 | -34.4032 | -16.8153
9.9255 | -34.4036 | -16.825
9.9316 | -34.404 | -16.8343
9.9377 | -34.4044 | -16.8437
9.9438 | -34.4048 | -16.8532
9.9499 | -34.4052 | -16.8629
9.956 | -34.4056 | -16.8723
9.9622 | -34.406 | -16.8818
9.9683 | -34.4064 | -16.8914
9.9744 | -34.4068 | -16.9009
9.9805 | -34.4072 | -16.9105
9.9866 | -34.4076 | -16.9197
9.9927 | -34.408 | -16.9294
9.9988 | -34.4084 | -16.9388
10.0049 | -34.4087 | -16.9483
10.011 | -34.4092 | -16.9577
10.0171 | -34.4096 | -16.9673
10.0232 | -34.41 | -16.9766
10.0293 | -34.4104 | -16.9859
10.0354 | -34.4109 | -16.9951
10.0415 | -34.4113 | -17.0045
10.0476 | -34.4117 | -17.0143
10.0537 | -34.412 | -17.024
10.0598 | -34.4124 | -17.0338
10.0659 | -34.4128 | -17.0433
10.072 | -34.4132 | -17.0528
10.0781 | -34.4137 | -17.0623
10.0842 | -34.414 | -17.0718
10.0903 | -34.4145 | -17.0817
10.0964 | -34.4149 | -17.0908
10.1026 | -34.4154 | -17.1003
10.1087 | -34.4158 | -17.1098
10.1148 | -34.4162 | -17.1194
10.1209 | -34.4166 | -17.1288
10.127 | -34.417 | -17.138
10.1331 | -34.4174 | -17.1473
10.1392 | -34.4178 | -17.1567
10.1453 | -34.4182 | -17.1663
10.1514 | -34.4186 | -17.1755
10.1575 | -34.419 | -17.1849
10.1636 | -34.4195 | -17.1936
10.1697 | -34.4198 | -17.2031
10.1758 | -34.4203 | -17.2123
10.1819 | -34.4207 | -17.2219
10.188 | -34.4211 | -17.2311
10.1941 | -34.4214 | -17.2406
10.2002 | -34.4218 | -17.2494
10.2063 | -34.4222 | -17.2591
10.2124 | -34.4226 | -17.2685
10.2185 | -34.423 | -17.2779
10.2246 | -34.4234 | -17.2872
10.2307 | -34.4238 | -17.2968
10.2368 | -34.4242 | -17.3054
10.2429 | -34.4247 | -17.3145
10.2491 | -34.4251 | -17.3236
10.2552 | -34.4255 | -17.3334
10.2613 | -34.4259 | -17.343
10.2674 | -34.4262 | -17.3526
10.2735 | -34.4266 | -17.3622
10.2796 | -34.427 | -17.3717
10.2857 | -34.4274 | -17.3811
10.2918 | -34.4278 | -17.3906
10.2979 | -34.4282 | -17.4002
10.304 | -34.4287 | -17.4099
10.3101 | -34.4291 | -17.4192
10.3162 | -34.4295 | -17.4286
10.3223 | -34.4299 | -17.438
10.3284 | -34.4303 | -17.4473
10.3345 | -34.4307 | -17.4564
10.3406 | -34.4311 | -17.4656
10.3467 | -34.4315 | -17.4749
10.3528 | -34.4319 | -17.484
10.3589 | -34.4323 | -17.4935
10.365 | -34.4327 | -17.5031
10.3711 | -34.4331 | -17.5128
10.3772 | -34.4334 | -17.5222
10.3833 | -34.4338 | -17.5315
10.3895 | -34.4342 | -17.5407
10.3956 | -34.4346 | -17.5498
10.4017 | -34.435 | -17.559
10.4078 | -34.4354 | -17.5684
10.4139 | -34.4359 | -17.5777
10.42 | -34.4363 | -17.5871
10.4261 | -34.4367 | -17.5963
10.4322 | -34.4371 | -17.6055
10.4383 | -34.4375 | -17.6149
10.4444 | -34.4379 | -17.6238
10.4505 | -34.4383 | -17.6331
10.4566 | -34.4387 | -17.6423
10.4627 | -34.4391 | -17.6516
10.4688 | -34.4396 | -17.6605
10.4749 | -34.4399 | -17.6696
10.481 | -34.4403 | -17.6789
10.4871 | -34.4407 | -17.6882
10.4932 | -34.4412 | -17.6974
10.4993 | -34.4416 | -17.7066
10.5054 | -34.442 | -17.716
10.5115 | -34.4424 | -17.7252
10.5176 | -34.4429 | -17.7343
10.5237 | -34.4433 | -17.7435
10.5298 | -34.4437 | -17.7527
10.536 | -34.4441 | -17.7623
10.5421 | -34.4445 | -17.7711
10.5482 | -34.4449 | -17.7806
10.5543 | -34.4454 | -17.7897
10.5604 | -34.4458 | -17.7986
10.5665 | -34.4462 | -17.8078
10.5726 | -34.4466 | -17.8167
10.5787 | -34.447 | -17.8262
10.5848 | -34.4475 | -17.835
10.5909 | -34.4478 | -17.8443
10.597 | -34.4482 | -17.8536
10.6031 | -34.4487 | -17.863
10.6092 | -34.4491 | -17.8725
10.6153 | -34.4494 | -17.8819
10.6214 | -34.4499 | -17.891
10.6275 | -34.4503 | -17.9002
10.6336 | -34.4507 | -17.9098
10.6397 | -34.4511 | -17.9187
10.6458 | -34.4515 | -17.9277
10.6519 | -34.4519 | -17.9368
10.658 | -34.4523 | -17.9459
10.6641 | -34.4527 | -17.9552
10.6702 | -34.4531 | -17.9647
10.6764 | -34.4536 | -17.9738
10.6825 | -34.454 | -17.983
10.6886 | -34.4544 | -17.9925
10.6947 | -34.4548 | -18.0014
10.7008 | -34.4552 | -18.0106
10.7069 | -34.4557 | -18.0198
10.713 | -34.4561 | -18.029
10.7191 | -34.4565 | -18.0383
10.7252 | -34.457 | -18.0471
10.7313 | -34.4574 | -18.0562
10.7374 | -34.4578 | -18.065
10.7435 | -34.4582 | -18.0746
10.7496 | -34.4586 | -18.0836
10.7557 | -34.459 | -18.0929
10.7618 | -34.4595 | -18.1022
10.7679 | -34.4599 | -18.1114
10.774 | -34.4602 | -18.1206
10.7801 | -34.4607 | -18.1301
10.7862 | -34.4611 | -18.1391
10.7923 | -34.4616 | -18.1482
10.7984 | -34.462 | -18.1571
10.8045 | -34.4624 | -18.1665
10.8106 | -34.4628 | -18.1757
10.8168 | -34.4632 | -18.1849
10.8229 | -34.4636 | -18.1941
10.829 | -34.464 | -18.2033
10.8351 | -34.4645 | -18.2125
10.8412 | -34.4648 | -18.2216
10.8473 | -34.4652 | -18.2303
10.8534 | -34.4657 | -18.2396
10.8595 | -34.4661 | -18.2485
10.8656 | -34.4665 | -18.2574
10.8717 | -34.4669 | -18.2662
10.8778 | -34.4674 | -18.2751
10.8839 | -34.4677 | -18.2841
10.89 | -34.4681 | -18.2931
10.8961 | -34.4686 | -18.3023
10.9022 | -34.469 | -18.3113
10.9083 | -34.4694 | -18.3207
10.9144 | -34.4699 | -18.3296
10.9205 | -34.4703 | -18.3387
10.9266 | -34.4708 | -18.3481
10.9327 | -34.4712 | -18.3572
10.9388 | -34.4716 | -18.3664
10.9449 | -34.472 | -18.3755
10.951 | -34.4725 | -18.3845
10.9571 | -34.4729 | -18.3937
10.9633 | -34.4733 | -18.4031
10.9694 | -34.4737 | -18.4125
10.9755 | -34.4742 | -18.4215
10.9816 | -34.4746 | -18.4308
10.9877 | -34.475 | -18.4399
10.9938 | -34.4754 | -18.4492
10.9999 | -34.4759 | -18.4583
11.006 | -34.4764 | -18.4673
11.0121 | -34.4767 | -18.4767
11.0182 | -34.4772 | -18.4854
11.0243 | -34.4776 | -18.4944
11.0304 | -34.478 | -18.5035
11.0365 | -34.4784 | -18.5125
11.0426 | -34.4789 | -18.5216
11.0487 | -34.4794 | -18.5306
11.0548 | -34.4799 | -18.5398
11.0609 | -34.4803 | -18.5493
11.067 | -34.4807 | -18.5584
11.0731 | -34.4811 | -18.5673
11.0792 | -34.4816 | -18.576
11.0853 | -34.482 | -18.5848
11.0914 | -34.4824 | -18.5938
11.0975 | -34.4829 | -18.6032
11.1037 | -34.4833 | -18.6121
11.1098 | -34.4836 | -18.6213
11.1159 | -34.4841 | -18.6305
11.122 | -34.4846 | -18.6397
11.1281 | -34.4849 | -18.6485
11.1342 | -34.4854 | -18.6573
11.1403 | -34.4858 | -18.6666
11.1464 | -34.4863 | -18.6756
11.1525 | -34.4867 | -18.6848
11.1586 | -34.4871 | -18.6943
11.1647 | -34.4875 | -18.7032
11.1708 | -34.4879 | -18.712
11.1769 | -34.4883 | -18.7205
11.183 | -34.4888 | -18.7294
11.1891 | -34.4893 | -18.7388
11.1952 | -34.4897 | -18.7482
11.2013 | -34.4901 | -18.7571
11.2074 | -34.4905 | -18.7659
11.2135 | -34.4908 | -18.775
11.2196 | -34.4912 | -18.7838
11.2257 | -34.4916 | -18.7926
11.2318 | -34.492 | -18.802
11.2379 | -34.4924 | -18.8108
11.244 | -34.4928 | -18.8198
11.2502 | -34.4932 | -18.8288
11.2563 | -34.4936 | -18.838
11.2624 | -34.494 | -18.8471
11.2685 | -34.4945 | -18.8559
11.2746 | -34.4948 | -18.8646
11.2807 | -34.4952 | -18.874
11.2868 | -34.4956 | -18.8834
11.2929 | -34.4961 | -18.8921
11.299 | -34.4965 | -18.9012
11.3051 | -34.4969 | -18.9101
11.3112 | -34.4973 | -18.9193
11.3173 | -34.4978 | -18.9284
11.3234 | -34.4982 | -18.9375
11.3295 | -34.4986 | -18.9465
11.3356 | -34.4991 | -18.9554
11.3417 | -34.4995 | -18.9644
11.3478 | -34.4999 | -18.9733
11.3539 | -34.5003 | -18.9822
11.36 | -34.5007 | -18.991
11.3661 | -34.5011 | -19
11.3722 | -34.5016 | -19.0088
11.3783 | -34.502 | -19.0175
11.3844 | -34.5024 | -19.0264
11.3906 | -34.5028 | -19.0357
11.3967 | -34.5032 | -19.0444
11.4028 | -34.5036 | -19.0535
11.4089 | -34.504 | -19.0624
11.415 | -34.5044 | -19.0713
11.4211 | -34.5049 | -19.0801
11.4272 | -34.5052 | -19.0892
11.4333 | -34.5057 | -19.0982
11.4394 | -34.5061 | -19.1077
11.4455 | -34.5065 | -19.1167
11.4516 | -34.5069 | -19.1254
11.4577 | -34.5074 | -19.1346
11.4638 | -34.5079 | -19.143
11.4699 | -34.5082 | -19.1523
11.476 | -34.5086 | -19.1618
11.4821 | -34.5091 | -19.1705
11.4882 | -34.5095 | -19.1799
11.4943 | -34.5099 | -19.1888
11.5004 | -34.5103 | -19.1978
11.5065 | -34.5108 | -19.2068
11.5126 | -34.5112 | -19.2154
11.5187 | -34.5117 | -19.2243
11.5248 | -34.5121 | -19.2334
11.5309 | -34.5125 | -19.2424
11.5371 | -34.5129 | -19.2513
11.5432 | -34.5132 | -19.2604
11.5493 | -34.5137 | -19.2696
11.5554 | -34.5141 | -19.2783
11.5615 | -34.5145 | -19.2875
11.5676 | -34.515 | -19.2967
11.5737 | -34.5154 | -19.3054
11.5798 | -34.5159 | -19.3141
11.5859 | -34.5163 | -19.323
11.592 | -34.5167 | -19.3319
11.5981 | -34.5171 | -19.3408
11.6042 | -34.5176 | -19.35
11.6103 | -34.518 | -19.3588
11.6164 | -34.5184 | -19.3675
11.6225 | -34.5188 | -19.3764
11.6286 | -34.5192 | -19.3855
11.6347 | -34.5196 | -19.3947
11.6408 | -34.5201 | -19.4035
11.6469 | -34.5205 | -19.4127
11.653 | -34.521 | -19.4215
11.6591 | -34.5214 | -19.4304
11.6652 | -34.5218 | -19.4394
11.6713 | -34.5222 | -19.4483
11.6775 | -34.5227 | -19.4571
11.6836 | -34.5231 | -19.4662
11.6897 | -34.5235 | -19.4749
11.6958 | -34.5239 | -19.4842
11.7019 | -34.5244 | -19.4932
11.708 | -34.5248 | -19.5023
11.7141 | -34.5253 | -19.5112
11.7202 | -34.5257 | -19.5198
11.7263 | -34.5261 | -19.5288
11.7324 | -34.5266 | -19.5379
11.7385 | -34.527 | -19.5469
11.7446 | -34.5274 | -19.556
11.7507 | -34.5279 | -19.5649
11.7568 | -34.5283 | -19.5738
11.7629 | -34.5287 | -19.5828
11.769 | -34.5291 | -19.5918
11.7751 | -34.5296 | -19.6004
11.7812 | -34.53 | -19.6092
11.7873 | -34.5304 | -19.6184
11.7934 | -34.5309 | -19.6273
11.7995 | -34.5313 | -19.6359
11.8056 | -34.5317 | -19.6452
11.8117 | -34.5321 | -19.6542
11.8178 | -34.5326 | -19.663
11.824 | -34.5331 | -19.6717
11.8301 | -34.5335 | -19.6802
11.8362 | -34.5339 | -19.6888
11.8423 | -34.5343 | -19.6978
11.8484 | -34.5348 | -19.7066
11.8545 | -34.5352 | -19.7158
11.8606 | -34.5356 | -19.7248
11.8667 | -34.5361 | -19.7337
11.8728 | -34.5365 | -19.7426
11.8789 | -34.5369 | -19.7519
11.885 | -34.5372 | -19.7607
11.8911 | -34.5377 | -19.7699
11.8972 | -34.5381 | -19.7785
11.9033 | -34.5385 | -19.7869
11.9094 | -34.5389 | -19.7954
11.9155 | -34.5393 | -19.8044
11.9216 | -34.5398 | -19.8132
11.9277 | -34.5402 | -19.8218
11.9338 | -34.5406 | -19.8307
11.9399 | -34.5411 | -19.8399
11.946 | -34.5415 | -19.8488
11.9521 | -34.5419 | -19.8576
11.9582 | -34.5424 | -19.8665
11.9644 | -34.5428 | -19.8754
11.9705 | -34.5433 | -19.8844
11.9766 | -34.5437 | -19.8933
11.9827 | -34.5441 | -19.9021
11.9888 | -34.5445 | -19.9106
11.9949 | -34.5449 | -19.9194
12.001 | -34.5454 | -19.9284
12.0071 | -34.5459 | -19.937
12.0132 | -34.5463 | -19.9455
12.0193 | -34.5467 | -19.9546
12.0254 | -34.5471 | -19.9633
12.0315 | -34.5476 | -19.9721
12.0376 | -34.548 | -19.9807
12.0437 | -34.5484 | -19.9893
12.0498 | -34.5488 | -19.9982
12.0559 | -34.5493 | -20.0071
12.062 | -34.5497 | -20.0159
12.0681 | -34.5501 | -20.0246
12.0742 | -34.5506 | -20.0337
12.0803 | -34.5511 | -20.0425
12.0864 | -34.5515 | -20.0514
12.0925 | -34.552 | -20.0602
12.0986 | -34.5524 | -20.0689
12.1047 | -34.5529 | -20.0777
12.1109 | -34.5533 | -20.0866
12.117 | -34.5538 | -20.0952
12.1231 | -34.5542 | -20.1041
12.1292 | -34.5546 | -20.1129
12.1353 | -34.5551 | -20.1214
12.1414 | -34.5556 | -20.1303
12.1475 | -34.556 | -20.139
12.1536 | -34.5564 | -20.1479
12.1597 | -34.5568 | -20.1569
12.1658 | -34.5573 | -20.1657
12.1719 | -34.5578 | -20.1745
12.178 | -34.5582 | -20.1835
12.1841 | -34.5586 | -20.1925
12.1902 | -34.5591 | -20.2014
12.1963 | -34.5595 | -20.2102
12.2024 | -34.56 | -20.2188
12.2085 | -34.5604 | -20.2276
12.2146 | -34.5608 | -20.2361
12.2207 | -34.5613 | -20.245
12.2268 | -34.5617 | -20.2538
12.2329 | -34.5621 | -20.2625
12.239 | -34.5625 | -20.2711
12.2451 | -34.563 | -20.2795
12.2513 | -34.5635 | -20.2883
12.2574 | -34.5639 | -20.2973
12.2635 | -34.5644 | -20.3057
12.2696 | -34.5648 | -20.3143
12.2757 | -34.5652 | -20.3226
12.2818 | -34.5656 | -20.3311
12.2879 | -34.566 | -20.3398
12.294 | -34.5665 | -20.3485
12.3001 | -34.5669 | -20.3574
12.3062 | -34.5674 | -20.3661
12.3123 | -34.5679 | -20.3748
12.3184 | -34.5683 | -20.3834
12.3245 | -34.5688 | -20.3923
12.3306 | -34.5692 | -20.4007
12.3367 | -34.5696 | -20.4092
12.3428 | -34.5701 | -20.4177
12.3489 | -34.5705 | -20.4259
12.355 | -34.571 | -20.4346
12.3611 | -34.5714 | -20.4433
12.3672 | -34.5719 | -20.4518
12.3733 | -34.5723 | -20.4609
12.3794 | -34.5727 | -20.4691
12.3855 | -34.5731 | -20.4776
12.3916 | -34.5735 | -20.4865
12.3978 | -34.574 | -20.4955
12.4039 | -34.5744 | -20.5042
12.41 | -34.5748 | -20.513
12.4161 | -34.5753 | -20.5214
12.4222 | -34.5758 | -20.5298
12.4283 | -34.5762 | -20.5385
12.4344 | -34.5767 | -20.5472
12.4405 | -34.5771 | -20.556
12.4466 | -34.5775 | -20.5646
12.4527 | -34.578 | -20.5732
12.4588 | -34.5784 | -20.5818
12.4649 | -34.5789 | -20.5908
12.471 | -34.5793 | -20.6
12.4771 | -34.5797 | -20.6081
12.4832 | -34.5802 | -20.6166
12.4893 | -34.5806 | -20.625
12.4954 | -34.5811 | -20.6335
12.5015 | -34.5816 | -20.6419
12.5076 | -34.582 | -20.6507
12.5137 | -34.5824 | -20.6597
12.5198 | -34.5828 | -20.668
12.5259 | -34.5833 | -20.6764
12.532 | -34.5838 | -20.6853
12.5382 | -34.5842 | -20.6938
12.5443 | -34.5846 | -20.7025
12.5504 | -34.585 | -20.7112
12.5565 | -34.5855 | -20.7196
12.5626 | -34.586 | -20.7284
12.5687 | -34.5864 | -20.7371
12.5748 | -34.5868 | -20.7457
12.5809 | -34.5873 | -20.7539
12.587 | -34.5877 | -20.763
12.5931 | -34.5881 | -20.7716
12.5992 | -34.5885 | -20.7805
12.6053 | -34.589 | -20.7889
12.6114 | -34.5895 | -20.7976
12.6175 | -34.5899 | -20.8063
12.6236 | -34.5903 | -20.8147
12.6297 | -34.5908 | -20.823
12.6358 | -34.5913 | -20.8314
12.6419 | -34.5917 | -20.8404
12.648 | -34.5921 | -20.8492
12.6541 | -34.5926 | -20.8578
12.6602 | -34.593 | -20.8659
12.6663 | -34.5934 | -20.8744
12.6724 | -34.5939 | -20.8832
12.6785 | -34.5943 | -20.8921
12.6847 | -34.5948 | -20.9007
12.6908 | -34.5952 | -20.9092
12.6969 | -34.5957 | -20.9178
12.703 | -34.5961 | -20.9265
12.7091 | -34.5966 | -20.9351
12.7152 | -34.597 | -20.9436
12.7213 | -34.5974 | -20.9523
12.7274 | -34.5979 | -20.9609
12.7335 | -34.5983 | -20.969
12.7396 | -34.5987 | -20.9771
12.7457 | -34.5992 | -20.9857
12.7518 | -34.5996 | -20.9943
12.7579 | -34.6 | -21.0032
12.764 | -34.6004 | -21.0118
12.7701 | -34.6007 | -21.02
12.7762 | -34.6012 | -21.0287
12.7823 | -34.6016 | -21.0373
12.7884 | -34.602 | -21.0458
12.7945 | -34.6025 | -21.0541
12.8006 | -34.6029 | -21.0619
12.8067 | -34.6034 | -21.0707
12.8128 | -34.6038 | -21.0788
12.8189 | -34.6041 | -21.0871
12.8251 | -34.6045 | -21.0954
12.8312 | -34.6049 | -21.104
12.8373 | -34.6054 | -21.1125
12.8434 | -34.6058 | -21.121
12.8495 | -34.6062 | -21.1294
12.8556 | -34.6066 | -21.1378
12.8617 | -34.6071 | -21.1461
12.8678 | -34.6075 | -21.1547
12.8739 | -34.6079 | -21.1627
12.88 | -34.6083 | -21.1714
12.8861 | -34.6087 | -21.18
12.8922 | -34.6091 | -21.1884
12.8983 | -34.6096 | -21.1965
12.9044 | -34.6101 | -21.2049
12.9105 | -34.6105 | -21.2134
12.9166 | -34.611 | -21.2222
12.9227 | -34.6114 | -21.231
12.9288 | -34.6119 | -21.2396
12.9349 | -34.6123 | -21.2483
12.941 | -34.6128 | -21.2569
12.9471 | -34.6132 | -21.2653
12.9532 | -34.6137 | -21.2739
12.9593 | -34.6141 | -21.2827
12.9654 | -34.6145 | -21.2914
12.9716 | -34.6149 | -21.2999
12.9777 | -34.6153 | -21.3081
12.9838 | -34.6157 | -21.3165
12.9899 | -34.6161 | -21.325
12.996 | -34.6165 | -21.3336
13.0021 | -34.6169 | -21.3421
13.0082 | -34.6174 | -21.3505
13.0143 | -34.6178 | -21.3586
13.0204 | -34.6181 | -21.3668
13.0265 | -34.6186 | -21.3751
13.0326 | -34.619 | -21.3839
13.0387 | -34.6195 | -21.3925
13.0448 | -34.6198 | -21.401
13.0509 | -34.6203 | -21.4097
13.057 | -34.6208 | -21.4178
13.0631 | -34.6213 | -21.426
13.0692 | -34.6217 | -21.4343
13.0753 | -34.6221 | -21.4427
13.0814 | -34.6225 | -21.4509
13.0875 | -34.623 | -21.459
13.0936 | -34.6234 | -21.4672
13.0997 | -34.6238 | -21.4752
13.1058 | -34.6242 | -21.483
13.112 | -34.6246 | -21.4914
13.1181 | -34.6251 | -21.4999
13.1242 | -34.6255 | -21.5084
13.1303 | -34.6259 | -21.5171
13.1364 | -34.6263 | -21.5255
13.1425 | -34.6267 | -21.5338
13.1486 | -34.6272 | -21.542
13.1547 | -34.6276 | -21.5501
13.1608 | -34.628 | -21.5584
13.1669 | -34.6285 | -21.5666
13.173 | -34.6289 | -21.5749
13.1791 | -34.6293 | -21.5833
13.1852 | -34.6297 | -21.592
13.1913 | -34.6301 | -21.6004
13.1974 | -34.6305 | -21.6088
13.2035 | -34.6309 | -21.6171
13.2096 | -34.6314 | -21.6254
13.2157 | -34.6318 | -21.6338
13.2218 | -34.6323 | -21.6422
13.2279 | -34.6327 | -21.6506
13.234 | -34.6331 | -21.6591
13.2401 | -34.6336 | -21.6673
13.2462 | -34.634 | -21.6756
13.2524 | -34.6344 | -21.6839
13.2585 | -34.6349 | -21.6924
13.2646 | -34.6352 | -21.7009
13.2707 | -34.6357 | -21.7094
13.2768 | -34.6361 | -21.7183
13.2829 | -34.6365 | -21.7269
13.289 | -34.6369 | -21.7353
13.2951 | -34.6373 | -21.7439
13.3012 | -34.6378 | -21.7525
13.3073 | -34.6382 | -21.7611
13.3134 | -34.6385 | -21.7696
13.3195 | -34.639 | -21.778
13.3256 | -34.6394 | -21.7867
13.3317 | -34.6398 | -21.7949
13.3378 | -34.6402 | -21.8036
13.3439 | -34.6406 | -21.8121
13.35 | -34.6411 | -21.8205
13.3561 | -34.6415 | -21.8288
13.3622 | -34.6419 | -21.8368
13.3683 | -34.6424 | -21.8449
13.3744 | -34.6427 | -21.8532
13.3805 | -34.6431 | -21.8617
13.3866 | -34.6435 | -21.8698
13.3927 | -34.6439 | -21.878
13.3989 | -34.6443 | -21.8862
13.405 | -34.6447 | -21.8943
13.4111 | -34.6451 | -21.9025
13.4172 | -34.6455 | -21.9105
13.4233 | -34.6459 | -21.919
13.4294 | -34.6463 | -21.927
13.4355 | -34.6468 | -21.9361
13.4416 | -34.6472 | -21.9444
13.4477 | -34.6475 | -21.9529
13.4538 | -34.6479 | -21.9613
13.4599 | -34.6484 | -21.9697
13.466 | -34.6487 | -21.9782
13.4721 | -34.6492 | -21.9863
13.4782 | -34.6496 | -21.9946
13.4843 | -34.6501 | -22.003
13.4904 | -34.6506 | -22.0111
13.4965 | -34.6509 | -22.02
13.5026 | -34.6513 | -22.0285
13.5087 | -34.6518 | -22.0365
13.5148 | -34.6521 | -22.0452
13.5209 | -34.6525 | -22.0533
13.527 | -34.653 | -22.0619
13.5331 | -34.6533 | -22.0699
13.5393 | -34.6538 | -22.0784
13.5454 | -34.6541 | -22.0868
13.5515 | -34.6546 | -22.0951
13.5576 | -34.655 | -22.1035
13.5637 | -34.6554 | -22.1118
13.5698 | -34.6558 | -22.1204
13.5759 | -34.6562 | -22.1286
13.582 | -34.6565 | -22.1369
13.5881 | -34.657 | -22.1453
13.5942 | -34.6574 | -22.1537
13.6003 | -34.6578 | -22.1621
13.6064 | -34.6582 | -22.1706
13.6125 | -34.6587 | -22.1794
13.6186 | -34.6591 | -22.1876
13.6247 | -34.6595 | -22.1959
13.6308 | -34.6599 | -22.204
13.6369 | -34.6603 | -22.2124
13.643 | -34.6608 | -22.2205
13.6491 | -34.6612 | -22.229
13.6552 | -34.6616 | -22.2372
13.6613 | -34.662 | -22.2457
13.6674 | -34.6624 | -22.2541
13.6735 | -34.6628 | -22.2625
13.6796 | -34.6633 | -22.2705
13.6858 | -34.6636 | -22.279
13.6919 | -34.6641 | -22.2874
13.698 | -34.6645 | -22.2957
13.7041 | -34.6648 | -22.3044
13.7102 | -34.6653 | -22.313
13.7163 | -34.6657 | -22.3216
13.7224 | -34.6662 | -22.3298
13.7285 | -34.6666 | -22.3383
13.7346 | -34.667 | -22.3468
13.7407 | -34.6674 | -22.3551
13.7468 | -34.6678 | -22.3636
13.7529 | -34.6682 | -22.3718
13.759 | -34.6686 | -22.3803
13.7651 | -34.669 | -22.3887
13.7712 | -34.6694 | -22.3973
13.7773 | -34.6698 | -22.4058
13.7834 | -34.6703 | -22.4142
13.7895 | -34.6708 | -22.4226
13.7956 | -34.6712 | -22.4311
13.8017 | -34.6716 | -22.4393
13.8078 | -34.672 | -22.4476
13.8139 | -34.6725 | -22.4557
13.82 | -34.6729 | -22.4638
13.8262 | -34.6733 | -22.4726
13.8323 | -34.6737 | -22.4811
13.8384 | -34.6741 | -22.4895
13.8445 | -34.6746 | -22.4982
13.8506 | -34.675 | -22.5065
13.8567 | -34.6754 | -22.5152
13.8628 | -34.6758 | -22.5237
13.8689 | -34.6762 | -22.5322
13.875 | -34.6767 | -22.5405
13.8811 | -34.6771 | -22.5488
13.8872 | -34.6776 | -22.5572
13.8933 | -34.678 | -22.5656
13.8994 | -34.6785 | -22.5737
13.9055 | -34.6789 | -22.582
13.9116 | -34.6793 | -22.5901
13.9177 | -34.6798 | -22.5982
13.9238 | -34.6802 | -22.6066
13.9299 | -34.6807 | -22.615
13.936 | -34.681 | -22.6232
13.9421 | -34.6815 | -22.6316
13.9482 | -34.6819 | -22.6401
13.9543 | -34.6823 | -22.6482
13.9604 | -34.6828 | -22.6567
13.9665 | -34.6832 | -22.6653
13.9727 | -34.6836 | -22.6735
13.9788 | -34.684 | -22.682
13.9849 | -34.6844 | -22.6905
13.991 | -34.6847 | -22.6991
13.9971 | -34.6851 | -22.7078
14.0032 | -34.6855 | -22.7165
14.0093 | -34.6859 | -22.7248
14.0154 | -34.6863 | -22.7336
14.0215 | -34.6867 | -22.7418
14.0276 | -34.6872 | -22.7499
14.0337 | -34.6876 | -22.7587
14.0398 | -34.688 | -22.7669
14.0459 | -34.6884 | -22.7755
14.052 | -34.6888 | -22.7838
14.0581 | -34.6892 | -22.7919
14.0642 | -34.6897 | -22.8004
14.0703 | -34.6901 | -22.8089
14.0764 | -34.6905 | -22.8176
14.0825 | -34.6909 | -22.8261
14.0886 | -34.6913 | -22.8341
14.0947 | -34.6917 | -22.8422
14.1008 | -34.6921 | -22.8509
14.1069 | -34.6926 | -22.8594
14.1131 | -34.6929 | -22.8681
14.1192 | -34.6933 | -22.8762
14.1253 | -34.6937 | -22.8842
14.1314 | -34.6942 | -22.8925
14.1375 | -34.6947 | -22.9011
14.1436 | -34.6951 | -22.9095
14.1497 | -34.6956 | -22.9176
14.1558 | -34.696 | -22.9262
14.1619 | -34.6964 | -22.9345
14.168 | -34.6968 | -22.9427
14.1741 | -34.6973 | -22.951
14.1802 | -34.6977 | -22.9593
14.1863 | -34.6981 | -22.9677
14.1924 | -34.6985 | -22.9761
14.1985 | -34.6989 | -22.9844
14.2046 | -34.6994 | -22.9927
14.2107 | -34.6998 | -23.0008
14.2168 | -34.7002 | -23.0096
14.2229 | -34.7007 | -23.0181
14.229 | -34.7011 | -23.0265
14.2351 | -34.7015 | -23.0345
14.2412 | -34.7019 | -23.0428
14.2473 | -34.7024 | -23.0512
14.2534 | -34.7029 | -23.0595
14.2596 | -34.7033 | -23.0679
14.2657 | -34.7037 | -23.0764
14.2718 | -34.7041 | -23.0849
14.2779 | -34.7045 | -23.0929
14.284 | -34.7049 | -23.1012
14.2901 | -34.7054 | -23.1093
14.2962 | -34.7059 | -23.1174
14.3023 | -34.7063 | -23.1258
14.3084 | -34.7067 | -23.1343
14.3145 | -34.7071 | -23.1427
14.3206 | -34.7075 | -23.1511
14.3267 | -34.7079 | -23.1594
14.3328 | -34.7084 | -23.1678
14.3389 | -34.7088 | -23.1759
14.345 | -34.7092 | -23.1844
14.3511 | -34.7096 | -23.1924
14.3572 | -34.71 | -23.201
14.3633 | -34.7104 | -23.2093
14.3694 | -34.7108 | -23.2176
14.3755 | -34.7112 | -23.2256
14.3816 | -34.7117 | -23.2337
14.3877 | -34.7121 | -23.2419
14.3938 | -34.7126 | -23.2505
14.4 | -34.713 | -23.2586
14.4061 | -34.7134 | -23.2669
14.4122 | -34.7138 | -23.2754
14.4183 | -34.7142 | -23.2837
14.4244 | -34.7146 | -23.2919
14.4305 | -34.715 | -23.3005
14.4366 | -34.7155 | -23.3089
14.4427 | -34.7159 | -23.3174
14.4488 | -34.7163 | -23.3254
14.4549 | -34.7168 | -23.3337
14.461 | -34.7172 | -23.3421
14.4671 | -34.7176 | -23.3506
14.4732 | -34.7181 | -23.3585
14.4793 | -34.7185 | -23.3675
14.4854 | -34.719 | -23.3754
14.4915 | -34.7194 | -23.3835
14.4976 | -34.7199 | -23.392
14.5037 | -34.7204 | -23.4004
14.5098 | -34.7208 | -23.4085
14.5159 | -34.7213 | -23.417
14.522 | -34.7217 | -23.4255
14.5281 | -34.7222 | -23.4336
14.5342 | -34.7225 | -23.442
14.5403 | -34.723 | -23.4503
14.5465 | -34.7235 | -23.459
14.5526 | -34.7239 | -23.4676
14.5587 | -34.7243 | -23.4758
14.5648 | -34.7248 | -23.484
14.5709 | -34.7252 | -23.4924
14.577 | -34.7256 | -23.5007
14.5831 | -34.726 | -23.5092
14.5892 | -34.7266 | -23.5172
14.5953 | -34.727 | -23.5251
14.6014 | -34.7275 | -23.5334
14.6075 | -34.7279 | -23.5415
14.6136 | -34.7283 | -23.5495
14.6197 | -34.7287 | -23.5579
14.6258 | -34.7292 | -23.5662
14.6319 | -34.7297 | -23.5743
14.638 | -34.7301 | -23.5825
14.6441 | -34.7306 | -23.5902
14.6502 | -34.731 | -23.5983
14.6563 | -34.7314 | -23.6064
14.6624 | -34.7319 | -23.6147
14.6685 | -34.7323 | -23.6227
14.6746 | -34.7328 | -23.631
14.6807 | -34.7332 | -23.6393
14.6869 | -34.7336 | -23.6474
14.693 | -34.7341 | -23.6556
14.6991 | -34.7345 | -23.6638
14.7052 | -34.735 | -23.6719
14.7113 | -34.7354 | -23.6804
14.7174 | -34.7358 | -23.6887
14.7235 | -34.7363 | -23.6967
14.7296 | -34.7367 | -23.7049
14.7357 | -34.7372 | -23.713
14.7418 | -34.7376 | -23.7209
14.7479 | -34.7381 | -23.7292
14.754 | -34.7385 | -23.7373
14.7601 | -34.739 | -23.7455
14.7662 | -34.7394 | -23.7537
14.7723 | -34.7399 | -23.7617
14.7784 | -34.7403 | -23.77
14.7845 | -34.7408 | -23.7782
14.7906 | -34.7412 | -23.7863
14.7967 | -34.7416 | -23.7944
14.8028 | -34.7421 | -23.8025
14.8089 | -34.7425 | -23.8105
14.815 | -34.7429 | -23.8184
14.8211 | -34.7433 | -23.8266
14.8272 | -34.7438 | -23.8349
14.8334 | -34.7442 | -23.8428
14.8395 | -34.7447 | -23.851
14.8456 | -34.7451 | -23.859
14.8517 | -34.7456 | -23.8673
14.8578 | -34.7461 | -23.8755
14.8639 | -34.7465 | -23.8839
14.87 | -34.747 | -23.8921
14.8761 | -34.7475 | -23.8999
14.8822 | -34.7479 | -23.9081
14.8883 | -34.7483 | -23.9162
14.8944 | -34.7488 | -23.9245
14.9005 | -34.7493 | -23.9324
14.9066 | -34.7497 | -23.9407
14.9127 | -34.7502 | -23.9487
14.9188 | -34.7506 | -23.9571
14.9249 | -34.7511 | -23.9652
14.931 | -34.7515 | -23.9732
14.9371 | -34.752 | -23.9812
14.9432 | -34.7524 | -23.9895
14.9493 | -34.7529 | -23.9975
14.9554 | -34.7533 | -24.0055
14.9615 | -34.7538 | -24.0136
14.9676 | -34.7542 | -24.0217
14.9738 | -34.7547 | -24.0295
14.9799 | -34.7551 | -24.038
14.986 | -34.7556 | -24.0456
14.9921 | -34.7561 | -24.0537
14.9982 | -34.7565 | -24.0621
15.0043 | -34.7569 | -24.0701
15.0104 | -34.7573 | -24.0781
15.0165 | -34.7578 | -24.0863
15.0226 | -34.7583 | -24.094
15.0287 | -34.7587 | -24.1025
15.0348 | -34.7592 | -24.1106
15.0409 | -34.7596 | -24.1189
15.047 | -34.7601 | -24.1273
15.0531 | -34.7607 | -24.1355
15.0592 | -34.7612 | -24.1438
15.0653 | -34.7616 | -24.1518
15.0714 | -34.762 | -24.1597
15.0775 | -34.7625 | -24.1674
15.0836 | -34.763 | -24.1755
15.0897 | -34.7635 | -24.1834
15.0958 | -34.764 | -24.1913
15.1019 | -34.7644 | -24.1993
15.108 | -34.7648 | -24.2074
15.1141 | -34.7653 | -24.2155
15.1203 | -34.7658 | -24.2236
15.1264 | -34.7663 | -24.2317
15.1325 | -34.7668 | -24.2398
15.1386 | -34.7672 | -24.2477
15.1447 | -34.7677 | -24.2556
15.1508 | -34.7682 | -24.2639
15.1569 | -34.7687 | -24.2722
15.163 | -34.7692 | -24.2798
15.1691 | -34.7697 | -24.2878
15.1752 | -34.7702 | -24.2958
15.1813 | -34.7706 | -24.3039
15.1874 | -34.7711 | -24.3118
15.1935 | -34.7716 | -24.3195
15.1996 | -34.7721 | -24.3277
15.2057 | -34.7726 | -24.3356
15.2118 | -34.773 | -24.3435
15.2179 | -34.7735 | -24.3513
15.224 | -34.7739 | -24.3593
15.2301 | -34.7744 | -24.3669
15.2362 | -34.7749 | -24.3748
15.2423 | -34.7753 | -24.3831
15.2484 | -34.7758 | -24.3911
15.2545 | -34.7763 | -24.399
15.2607 | -34.7767 | -24.4072
15.2668 | -34.7773 | -24.4148
15.2729 | -34.7778 | -24.4227
15.279 | -34.7782 | -24.4307
15.2851 | -34.7787 | -24.4387
15.2912 | -34.7792 | -24.4463
15.2973 | -34.7796 | -24.4543
15.3034 | -34.7802 | -24.462
15.3095 | -34.7806 | -24.4697
15.3156 | -34.7811 | -24.4778
15.3217 | -34.7815 | -24.4854
15.3278 | -34.782 | -24.4935
15.3339 | -34.7825 | -24.5012
15.34 | -34.783 | -24.509
15.3461 | -34.7834 | -24.517
15.3522 | -34.7839 | -24.525
15.3583 | -34.7844 | -24.5329
15.3644 | -34.7848 | -24.541
15.3705 | -34.7853 | -24.5485
15.3766 | -34.7858 | -24.5567
15.3827 | -34.7863 | -24.5643
15.3888 | -34.7868 | -24.572
15.3949 | -34.7872 | -24.5796
15.401 | -34.7877 | -24.5876
15.4072 | -34.7882 | -24.5952
15.4133 | -34.7887 | -24.6031
15.4194 | -34.7891 | -24.6112
15.4255 | -34.7896 | -24.6191
15.4316 | -34.7901 | -24.6271
15.4377 | -34.7906 | -24.635
15.4438 | -34.7911 | -24.6424
15.4499 | -34.7916 | -24.6503
15.456 | -34.792 | -24.658
15.4621 | -34.7925 | -24.6659
15.4682 | -34.793 | -24.674
15.4743 | -34.7935 | -24.6817
15.4804 | -34.794 | -24.6897
15.4865 | -34.7944 | -24.6975
15.4926 | -34.7949 | -24.7055
15.4987 | -34.7954 | -24.7135
15.5048 | -34.7959 | -24.7212
15.5109 | -34.7964 | -24.7287
15.517 | -34.7969 | -24.7367
15.5231 | -34.7974 | -24.7446
15.5292 | -34.7978 | -24.7522
15.5353 | -34.7983 | -24.7599
15.5414 | -34.7988 | -24.768
15.5476 | -34.7993 | -24.7758
15.5537 | -34.7998 | -24.7836
15.5598 | -34.8003 | -24.7911
15.5659 | -34.8008 | -24.7993
15.572 | -34.8013 | -24.8068
15.5781 | -34.8017 | -24.8147
15.5842 | -34.8023 | -24.8221
15.5903 | -34.8028 | -24.8295
15.5964 | -34.8033 | -24.8371
15.6025 | -34.8038 | -24.8448
15.6086 | -34.8042 | -24.8527
15.6147 | -34.8047 | -24.8604
15.6208 | -34.8052 | -24.868
15.6269 | -34.8057 | -24.8757
15.633 | -34.8062 | -24.8832
15.6391 | -34.8066 | -24.8912
15.6452 | -34.8071 | -24.8989
15.6513 | -34.8076 | -24.9068
15.6574 | -34.8081 | -24.9144
15.6635 | -34.8086 | -24.9221
15.6696 | -34.8091 | -24.9297
15.6757 | -34.8097 | -24.9378
15.6818 | -34.8101 | -24.9459
15.688 | -34.8107 | -24.9533
15.6941 | -34.8111 | -24.961
15.7002 | -34.8116 | -24.9685
15.7063 | -34.8121 | -24.9763
15.7124 | -34.8125 | -24.984
15.7185 | -34.813 | -24.9915
15.7246 | -34.8135 | -24.9991
15.7307 | -34.814 | -25.0071
15.7368 | -34.8145 | -25.0152
15.7429 | -34.815 | -25.0229
15.749 | -34.8154 | -25.0308
15.7551 | -34.8159 | -25.0385
15.7612 | -34.8164 | -25.0456
15.7673 | -34.8169 | -25.0533
15.7734 | -34.8173 | -25.0607
15.7795 | -34.8178 | -25.0687
15.7856 | -34.8183 | -25.0762
15.7917 | -34.8188 | -25.0836
15.7978 | -34.8194 | -25.0913
15.8039 | -34.8198 | -25.099
15.81 | -34.8203 | -25.1065
15.8161 | -34.8208 | -25.1144
15.8222 | -34.8213 | -25.1221
15.8283 | -34.8218 | -25.1294
15.8345 | -34.8223 | -25.1372
15.8406 | -34.8228 | -25.1447
15.8467 | -34.8233 | -25.1523
15.8528 | -34.8238 | -25.1598
15.8589 | -34.8243 | -25.1675
15.865 | -34.8247 | -25.1749
15.8711 | -34.8253 | -25.1827
15.8772 | -34.8258 | -25.1904
15.8833 | -34.8263 | -25.1985
15.8894 | -34.8267 | -25.206
15.8955 | -34.8272 | -25.2137
15.9016 | -34.8277 | -25.2215
15.9077 | -34.8282 | -25.2292
15.9138 | -34.8287 | -25.2368
15.9199 | -34.8292 | -25.2447
15.926 | -34.8297 | -25.2525
15.9321 | -34.8302 | -25.2601
15.9382 | -34.8308 | -25.2677
15.9443 | -34.8313 | -25.2756
15.9504 | -34.8318 | -25.2832
15.9565 | -34.8323 | -25.2909
15.9626 | -34.8328 | -25.2985
15.9687 | -34.8334 | -25.3058
15.9749 | -34.8339 | -25.3135
15.981 | -34.8344 | -25.3212
15.9871 | -34.8349 | -25.3288
15.9932 | -34.8354 | -25.3365
15.9993 | -34.8359 | -25.3439
16.0054 | -34.8363 | -25.3514
16.0115 | -34.8369 | -25.3589
16.0176 | -34.8374 | -25.3668
16.0237 | -34.8379 | -25.3746
16.0298 | -34.8384 | -25.3818
16.0359 | -34.8389 | -25.3893
16.042 | -34.8394 | -25.3968
16.0481 | -34.8399 | -25.4044
16.0542 | -34.8403 | -25.4124
16.0603 | -34.8409 | -25.4199
16.0664 | -34.8414 | -25.4275
16.0725 | -34.8419 | -25.4351
16.0786 | -34.8424 | -25.4427
16.0847 | -34.8429 | -25.4501
16.0908 | -34.8434 | -25.4576
16.0969 | -34.8439 | -25.4649
16.103 | -34.8444 | -25.4725
16.1091 | -34.8449 | -25.4799
16.1152 | -34.8453 | -25.4876
16.1214 | -34.8458 | -25.4954
16.1275 | -34.8463 | -25.5029
16.1336 | -34.8468 | -25.5104
16.1397 | -34.8473 | -25.518
16.1458 | -34.8479 | -25.5254
16.1519 | -34.8484 | -25.5331
16.158 | -34.8489 | -25.5407
16.1641 | -34.8494 | -25.5483
16.1702 | -34.8498 | -25.5562
16.1763 | -34.8503 | -25.5638
16.1824 | -34.8509 | -25.571
16.1885 | -34.8514 | -25.5785
16.1946 | -34.8519 | -25.5861
16.2007 | -34.8523 | -25.594
16.2068 | -34.8528 | -25.6014
16.2129 | -34.8533 | -25.6091
16.219 | -34.8538 | -25.6165
16.2251 | -34.8543 | -25.6243
16.2312 | -34.8547 | -25.6319
16.2373 | -34.8552 | -25.6394
16.2434 | -34.8557 | -25.6468
16.2495 | -34.8562 | -25.6541
16.2556 | -34.8567 | -25.6618
16.2618 | -34.8572 | -25.6689
16.2679 | -34.8578 | -25.6765
16.274 | -34.8582 | -25.6843
16.2801 | -34.8587 | -25.6916
16.2862 | -34.8592 | -25.6989
16.2923 | -34.8597 | -25.7068
16.2984 | -34.8602 | -25.7145
16.3045 | -34.8608 | -25.7219
16.3106 | -34.8613 | -25.7291
16.3167 | -34.8618 | -25.7367
16.3228 | -34.8623 | -25.7443
16.3289 | -34.8628 | -25.7518
16.335 | -34.8633 | -25.7596
16.3411 | -34.8637 | -25.7667
16.3472 | -34.8643 | -25.7744
16.3533 | -34.8648 | -25.782
16.3594 | -34.8652 | -25.7892
16.3655 | -34.8657 | -25.7969
16.3716 | -34.8662 | -25.8044
16.3777 | -34.8667 | -25.8119
16.3838 | -34.8672 | -25.8193
16.3899 | -34.8676 | -25.827
16.396 | -34.8681 | -25.8347
16.4021 | -34.8687 | -25.8422
16.4083 | -34.8692 | -25.8496
16.4144 | -34.8697 | -25.8569
16.4205 | -34.8702 | -25.8646
16.4266 | -34.8707 | -25.872
16.4327 | -34.8712 | -25.8793
16.4388 | -34.8717 | -25.8868
16.4449 | -34.8722 | -25.8942
16.451 | -34.8727 | -25.9019
16.4571 | -34.8731 | -25.9095
16.4632 | -34.8736 | -25.9166
16.4693 | -34.8741 | -25.9242
16.4754 | -34.8746 | -25.9319
16.4815 | -34.8751 | -25.9395
16.4876 | -34.8756 | -25.9468
16.4937 | -34.8761 | -25.9543
16.4998 | -34.8766 | -25.9615
16.5059 | -34.8771 | -25.969
16.512 | -34.8776 | -25.9769
16.5181 | -34.8782 | -25.9845
16.5242 | -34.8786 | -25.9921
16.5303 | -34.8791 | -25.9994
16.5364 | -34.8796 | -26.0068
16.5425 | -34.8801 | -26.0144
16.5487 | -34.8806 | -26.022
16.5548 | -34.8811 | -26.0294
16.5609 | -34.8816 | -26.0371
16.567 | -34.8821 | -26.0446
16.5731 | -34.8825 | -26.0523
16.5792 | -34.883 | -26.0598
16.5853 | -34.8836 | -26.0671
16.5914 | -34.8841 | -26.0747
16.5975 | -34.8846 | -26.0825
16.6036 | -34.885 | -26.0898
16.6097 | -34.8855 | -26.0974
16.6158 | -34.886 | -26.1051
16.6219 | -34.8865 | -26.1124
16.628 | -34.887 | -26.1197
16.6341 | -34.8875 | -26.127
16.6402 | -34.888 | -26.134
16.6463 | -34.8885 | -26.1415
16.6524 | -34.889 | -26.1489
16.6585 | -34.8895 | -26.1564
16.6646 | -34.89 | -26.164
16.6707 | -34.8905 | -26.1709
16.6768 | -34.891 | -26.1784
16.6829 | -34.8915 | -26.186
16.689 | -34.892 | -26.1935
16.6952 | -34.8925 | -26.201
16.7013 | -34.893 | -26.2086
16.7074 | -34.8935 | -26.2159
16.7135 | -34.894 | -26.2234
16.7196 | -34.8945 | -26.2307
16.7257 | -34.895 | -26.2379
16.7318 | -34.8954 | -26.2454
16.7379 | -34.8959 | -26.253
16.744 | -34.8965 | -26.2602
16.7501 | -34.897 | -26.2676
16.7562 | -34.8975 | -26.2749
16.7623 | -34.898 | -26.2823
16.7684 | -34.8985 | -26.2898
16.7745 | -34.899 | -26.297
16.7806 | -34.8995 | -26.3047
16.7867 | -34.8999 | -26.312
16.7928 | -34.9004 | -26.3196
16.7989 | -34.9009 | -26.3271
16.805 | -34.9013 | -26.3346
16.8111 | -34.9018 | -26.3421
16.8172 | -34.9023 | -26.3499
16.8233 | -34.9028 | -26.3575
16.8294 | -34.9033 | -26.3648
16.8356 | -34.9038 | -26.3727
16.8417 | -34.9043 | -26.3801
16.8478 | -34.9048 | -26.3874
16.8539 | -34.9052 | -26.3947
16.86 | -34.9057 | -26.4019
16.8661 | -34.9062 | -26.4094
16.8722 | -34.9067 | -26.4167
16.8783 | -34.9072 | -26.424
16.8844 | -34.9077 | -26.4314
16.8905 | -34.9082 | -26.439
16.8966 | -34.9087 | -26.4465
16.9027 | -34.9092 | -26.4536
16.9088 | -34.9097 | -26.4615
16.9149 | -34.9102 | -26.4694
16.921 | -34.9107 | -26.4766
16.9271 | -34.9112 | -26.4844
16.9332 | -34.9117 | -26.4918
16.9393 | -34.9121 | -26.4995
16.9454 | -34.9126 | -26.5066
16.9515 | -34.913 | -26.5143
16.9576 | -34.9135 | -26.5218
16.9637 | -34.914 | -26.5294
16.9698 | -34.9145 | -26.5365
16.9759 | -34.915 | -26.5439
16.9821 | -34.9155 | -26.5515
16.9882 | -34.916 | -26.559
16.9943 | -34.9165 | -26.5665
17.0004 | -34.917 | -26.5738
17.0065 | -34.9175 | -26.5811
17.0126 | -34.9179 | -26.5885
17.0187 | -34.9184 | -26.596
17.0248 | -34.9189 | -26.6032
17.0309 | -34.9194 | -26.6107
17.037 | -34.9199 | -26.618
17.0431 | -34.9203 | -26.6251
17.0492 | -34.9208 | -26.6325
17.0553 | -34.9212 | -26.64
17.0614 | -34.9217 | -26.6473
17.0675 | -34.9222 | -26.655
17.0736 | -34.9227 | -26.6627
17.0797 | -34.9232 | -26.6696
17.0858 | -34.9237 | -26.6771
17.0919 | -34.9241 | -26.6847
17.098 | -34.9246 | -26.6921
17.1041 | -34.9251 | -26.6992
17.1102 | -34.9256 | -26.7069
17.1163 | -34.9261 | -26.7144
17.1225 | -34.9266 | -26.7217
17.1286 | -34.9271 | -26.7287
17.1347 | -34.9275 | -26.7361
17.1408 | -34.928 | -26.7434
17.1469 | -34.9286 | -26.7508
17.153 | -34.9291 | -26.7581
17.1591 | -34.9295 | -26.7653
17.1652 | -34.93 | -26.7726
17.1713 | -34.9305 | -26.7799
17.1774 | -34.9309 | -26.7874
17.1835 | -34.9315 | -26.7951
17.1896 | -34.9319 | -26.8024
17.1957 | -34.9324 | -26.8099
17.2018 | -34.9328 | -26.8175
17.2079 | -34.9333 | -26.825
17.214 | -34.9338 | -26.8322
17.2201 | -34.9343 | -26.8397
17.2262 | -34.9347 | -26.8472
17.2323 | -34.9352 | -26.8546
17.2384 | -34.9357 | -26.8619
17.2445 | -34.9362 | -26.8694
17.2506 | -34.9366 | -26.8768
17.2567 | -34.9371 | -26.8844
17.2628 | -34.9376 | -26.8922
17.269 | -34.938 | -26.8996
17.2751 | -34.9385 | -26.9071
17.2812 | -34.939 | -26.9144
17.2873 | -34.9395 | -26.922
17.2934 | -34.9399 | -26.9293
17.2995 | -34.9404 | -26.9367
17.3056 | -34.9409 | -26.9439
17.3117 | -34.9413 | -26.9512
17.3178 | -34.9418 | -26.9584
17.3239 | -34.9423 | -26.9658
17.33 | -34.9428 | -26.973
17.3361 | -34.9432 | -26.9805
17.3422 | -34.9437 | -26.988
17.3483 | -34.9442 | -26.9956
17.3544 | -34.9446 | -27.0031
17.3605 | -34.9451 | -27.0108
17.3666 | -34.9456 | -27.0184
17.3727 | -34.9461 | -27.0255
17.3788 | -34.9465 | -27.0325
17.3849 | -34.947 | -27.0398
17.391 | -34.9475 | -27.0469
17.3971 | -34.948 | -27.0544
17.4032 | -34.9484 | -27.0622
17.4094 | -34.9489 | -27.0693
17.4155 | -34.9494 | -27.0765
17.4216 | -34.9498 | -27.0838
17.4277 | -34.9503 | -27.0915
17.4338 | -34.9508 | -27.0988
17.4399 | -34.9513 | -27.1063
17.446 | -34.9517 | -27.1135
17.4521 | -34.9522 | -27.1207
17.4582 | -34.9527 | -27.1282
17.4643 | -34.9531 | -27.1356
17.4704 | -34.9536 | -27.1429
17.4765 | -34.9541 | -27.1505
17.4826 | -34.9546 | -27.1579
17.4887 | -34.955 | -27.1656
17.4948 | -34.9555 | -27.1728
17.5009 | -34.9559 | -27.1799
17.507 | -34.9564 | -27.1874
17.5131 | -34.9569 | -27.1947
17.5192 | -34.9573 | -27.2021
17.5253 | -34.9578 | -27.2096
17.5314 | -34.9582 | -27.217
17.5375 | -34.9587 | -27.2244
17.5436 | -34.9591 | -27.2318
17.5497 | -34.9596 | -27.2393
17.5559 | -34.96 | -27.2464
17.562 | -34.9605 | -27.2536
17.5681 | -34.961 | -27.2609
17.5742 | -34.9614 | -27.2683
17.5803 | -34.9619 | -27.2753
17.5864 | -34.9624 | -27.2824
17.5925 | -34.9629 | -27.2901
17.5986 | -34.9634 | -27.2972
17.6047 | -34.9639 | -27.3048
17.6108 | -34.9643 | -27.3118
17.6169 | -34.9648 | -27.3191
17.623 | -34.9653 | -27.3264
17.6291 | -34.9658 | -27.3337
17.6352 | -34.9662 | -27.3409
17.6413 | -34.9667 | -27.3483
17.6474 | -34.9671 | -27.356
17.6535 | -34.9676 | -27.3634
17.6596 | -34.968 | -27.3707
17.6657 | -34.9685 | -27.378
17.6718 | -34.9689 | -27.3856
17.6779 | -34.9694 | -27.3928
17.684 | -34.9699 | -27.4001
17.6901 | -34.9704 | -27.4075
17.6963 | -34.9708 | -27.4146
17.7024 | -34.9713 | -27.4218
17.7085 | -34.9717 | -27.429
17.7146 | -34.9722 | -27.4363
17.7207 | -34.9727 | -27.4434
17.7268 | -34.9732 | -27.4507
17.7329 | -34.9736 | -27.4581
17.739 | -34.9741 | -27.4657
17.7451 | -34.9746 | -27.4731
17.7512 | -34.9751 | -27.4804
17.7573 | -34.9755 | -27.4877
17.7634 | -34.976 | -27.4948
17.7695 | -34.9764 | -27.5021
17.7756 | -34.9769 | -27.5094
17.7817 | -34.9774 | -27.5165
17.7878 | -34.9779 | -27.524
17.7939 | -34.9783 | -27.5312
17.8 | -34.9788 | -27.5383
17.8061 | -34.9793 | -27.5459
17.8122 | -34.9797 | -27.5531
17.8183 | -34.9802 | -27.5605
17.8244 | -34.9807 | -27.568
17.8305 | -34.9812 | -27.5752
17.8366 | -34.9816 | -27.5822
17.8428 | -34.9821 | -27.5897
17.8489 | -34.9826 | -27.5969
17.855 | -34.983 | -27.604
17.8611 | -34.9835 | -27.6111
17.8672 | -34.9839 | -27.6182
17.8733 | -34.9844 | -27.6257
17.8794 | -34.9849 | -27.6331
17.8855 | -34.9854 | -27.6404
17.8916 | -34.9859 | -27.6472
17.8977 | -34.9864 | -27.6547
17.9038 | -34.9868 | -27.6619
17.9099 | -34.9873 | -27.6692
17.916 | -34.9878 | -27.6766
17.9221 | -34.9882 | -27.6837
17.9282 | -34.9887 | -27.691
17.9343 | -34.9892 | -27.6982
17.9404 | -34.9896 | -27.7056
17.9465 | -34.9901 | -27.7131
17.9526 | -34.9905 | -27.7206
17.9587 | -34.991 | -27.7279
17.9648 | -34.9915 | -27.735
17.9709 | -34.9919 | -27.7423
17.977 | -34.9924 | -27.7494
17.9832 | -34.9929 | -27.7567
17.9893 | -34.9933 | -27.7642
17.9954 | -34.9938 | -27.7718
18.0015 | -34.9943 | -27.7791
18.0076 | -34.9947 | -27.786
18.0137 | -34.9951 | -27.7935
18.0198 | -34.9956 | -27.8006
18.0259 | -34.996 | -27.808
18.032 | -34.9965 | -27.8151
18.0381 | -34.9969 | -27.8222
18.0442 | -34.9974 | -27.8296
18.0503 | -34.9979 | -27.837
18.0564 | -34.9983 | -27.8442
18.0625 | -34.9988 | -27.8516
18.0686 | -34.9993 | -27.8587
18.0747 | -34.9997 | -27.8659
18.0808 | -35.0002 | -27.8729
18.0869 | -35.0006 | -27.8803
18.093 | -35.0011 | -27.8874
18.0991 | -35.0016 | -27.8949
18.1052 | -35.0021 | -27.9025
18.1113 | -35.0025 | -27.9094
18.1174 | -35.003 | -27.9164
18.1236 | -35.0035 | -27.9237
18.1297 | -35.004 | -27.9305
18.1358 | -35.0045 | -27.9378
18.1419 | -35.005 | -27.9452
18.148 | -35.0055 | -27.9525
18.1541 | -35.006 | -27.9597
18.1602 | -35.0065 | -27.967
18.1663 | -35.007 | -27.9742
18.1724 | -35.0074 | -27.9816
18.1785 | -35.0079 | -27.9888
18.1846 | -35.0083 | -27.9959
18.1907 | -35.0088 | -28.0035
18.1968 | -35.0092 | -28.0109
18.2029 | -35.0096 | -28.0178
18.209 | -35.0101 | -28.0254
18.2151 | -35.0106 | -28.0329
18.2212 | -35.0111 | -28.0397
18.2273 | -35.0116 | -28.0471
18.2334 | -35.012 | -28.0544
18.2395 | -35.0125 | -28.0615
18.2456 | -35.0129 | -28.0686
18.2517 | -35.0134 | -28.0758
18.2578 | -35.0138 | -28.0827
18.2639 | -35.0143 | -28.0902
18.2701 | -35.0148 | -28.0974
18.2762 | -35.0153 | -28.1048
18.2823 | -35.0158 | -28.1121
18.2884 | -35.0162 | -28.1194
18.2945 | -35.0167 | -28.1264
18.3006 | -35.0171 | -28.1337
18.3067 | -35.0176 | -28.1408
18.3128 | -35.0181 | -28.148
18.3189 | -35.0185 | -28.1555
18.325 | -35.019 | -28.1623
18.3311 | -35.0194 | -28.1696
18.3372 | -35.0199 | -28.1765
18.3433 | -35.0203 | -28.1834
18.3494 | -35.0208 | -28.1906
18.3555 | -35.0212 | -28.1976
18.3616 | -35.0217 | -28.2046
18.3677 | -35.0222 | -28.2117
18.3738 | -35.0226 | -28.2188
18.3799 | -35.0231 | -28.2257
18.386 | -35.0236 | -28.2329
18.3921 | -35.0241 | -28.24
18.3982 | -35.0245 | -28.247
18.4043 | -35.0249 | -28.2542
18.4105 | -35.0253 | -28.2612
18.4166 | -35.0258 | -28.2682
18.4227 | -35.0263 | -28.2754
18.4288 | -35.0268 | -28.2827
18.4349 | -35.0273 | -28.2897
18.441 | -35.0278 | -28.2968
18.4471 | -35.0282 | -28.304
18.4532 | -35.0286 | -28.3112
18.4593 | -35.0291 | -28.3185
18.4654 | -35.0295 | -28.3255
18.4715 | -35.03 | -28.3328
18.4776 | -35.0305 | -28.3401
18.4837 | -35.031 | -28.3473
18.4898 | -35.0314 | -28.3546
18.4959 | -35.0319 | -28.3619
18.502 | -35.0324 | -28.3692
18.5081 | -35.0329 | -28.3764
18.5142 | -35.0333 | -28.3838
18.5203 | -35.0338 | -28.3906
18.5264 | -35.0342 | -28.3976
18.5325 | -35.0347 | -28.4048
18.5386 | -35.0352 | -28.4123
18.5447 | -35.0357 | -28.4192
18.5508 | -35.0362 | -28.4262
18.557 | -35.0366 | -28.4335
18.5631 | -35.0371 | -28.4408
18.5692 | -35.0375 | -28.4478
18.5753 | -35.038 | -28.4547
18.5814 | -35.0385 | -28.4619
18.5875 | -35.0389 | -28.4695
18.5936 | -35.0394 | -28.4769
18.5997 | -35.0398 | -28.4842
18.6058 | -35.0403 | -28.4914
18.6119 | -35.0407 | -28.4988
18.618 | -35.0412 | -28.5061
18.6241 | -35.0417 | -28.5134
18.6302 | -35.0422 | -28.5208
18.6363 | -35.0426 | -28.5277
18.6424 | -35.0431 | -28.5346
18.6485 | -35.0435 | -28.5418
18.6546 | -35.044 | -28.5488
18.6607 | -35.0444 | -28.5558
18.6668 | -35.0449 | -28.5629
18.6729 | -35.0454 | -28.5702
18.679 | -35.0459 | -28.577
18.6851 | -35.0463 | -28.5841
18.6912 | -35.0468 | -28.5912
18.6974 | -35.0472 | -28.5983
18.7035 | -35.0477 | -28.6054
18.7096 | -35.0482 | -28.6127
18.7157 | -35.0486 | -28.6199
18.7218 | -35.0491 | -28.627
18.7279 | -35.0496 | -28.6342
18.734 | -35.0501 | -28.6415
18.7401 | -35.0505 | -28.6484
18.7462 | -35.051 | -28.6549
18.7523 | -35.0515 | -28.662
18.7584 | -35.052 | -28.6691
18.7645 | -35.0524 | -28.6761
18.7706 | -35.0529 | -28.6833
18.7767 | -35.0533 | -28.6902
18.7828 | -35.0538 | -28.6969
18.7889 | -35.0543 | -28.7039
18.795 | -35.0547 | -28.7109
18.8011 | -35.0552 | -28.7181
18.8072 | -35.0557 | -28.7252
18.8133 | -35.0561 | -28.7324
18.8194 | -35.0566 | -28.7393
18.8255 | -35.0571 | -28.7461
18.8316 | -35.0575 | -28.7533
18.8377 | -35.0579 | -28.7602
18.8439 | -35.0584 | -28.7672
18.85 | -35.0589 | -28.7744
18.8561 | -35.0593 | -28.7816
18.8622 | -35.0598 | -28.7888
18.8683 | -35.0603 | -28.7956
18.8744 | -35.0607 | -28.8024
18.8805 | -35.0612 | -28.8096
18.8866 | -35.0617 | -28.8167
18.8927 | -35.0622 | -28.8235
18.8988 | -35.0627 | -28.8304
18.9049 | -35.0631 | -28.8373
18.911 | -35.0637 | -28.8444
18.9171 | -35.0642 | -28.8513
18.9232 | -35.0646 | -28.8582
18.9293 | -35.0651 | -28.8651
18.9354 | -35.0655 | -28.8719
18.9415 | -35.0659 | -28.8789
18.9476 | -35.0664 | -28.886
18.9537 | -35.0669 | -28.8932
18.9598 | -35.0675 | -28.8999
18.9659 | -35.068 | -28.9069
18.972 | -35.0684 | -28.9138
18.9781 | -35.0689 | -28.921
18.9843 | -35.0693 | -28.9283
18.9904 | -35.0697 | -28.9354
18.9965 | -35.0702 | -28.9425
19.0026 | -35.0707 | -28.9498
19.0087 | -35.0712 | -28.9567
19.0148 | -35.0717 | -28.9637
19.0209 | -35.0722 | -28.9711
19.027 | -35.0727 | -28.9784
19.0331 | -35.0732 | -28.9856
19.0392 | -35.0737 | -28.9926
19.0453 | -35.0742 | -28.9998
19.0514 | -35.0746 | -29.0069
19.0575 | -35.075 | -29.014
19.0636 | -35.0755 | -29.0205
19.0697 | -35.076 | -29.0279
19.0758 | -35.0765 | -29.0348
19.0819 | -35.0769 | -29.0421
19.088 | -35.0774 | -29.0488
19.0941 | -35.0779 | -29.0554
19.1002 | -35.0783 | -29.0622
19.1063 | -35.0788 | -29.0694
19.1124 | -35.0794 | -29.0761
19.1185 | -35.0798 | -29.0832
19.1246 | -35.0803 | -29.0901
19.1308 | -35.0808 | -29.097
19.1369 | -35.0813 | -29.1039
19.143 | -35.0817 | -29.1112
19.1491 | -35.0822 | -29.1183
19.1552 | -35.0827 | -29.1256
19.1613 | -35.0831 | -29.1324
19.1674 | -35.0836 | -29.1393
19.1735 | -35.084 | -29.1463
19.1796 | -35.0845 | -29.1531
19.1857 | -35.085 | -29.1601
19.1918 | -35.0855 | -29.1672
19.1979 | -35.0859 | -29.1743
19.204 | -35.0864 | -29.181
19.2101 | -35.0869 | -29.188
19.2162 | -35.0873 | -29.1952
19.2223 | -35.0877 | -29.202
19.2284 | -35.0882 | -29.2095
19.2345 | -35.0887 | -29.2164
19.2406 | -35.0892 | -29.2235
19.2467 | -35.0897 | -29.2306
19.2528 | -35.0901 | -29.2376
19.2589 | -35.0906 | -29.2445
19.265 | -35.0911 | -29.2514
19.2712 | -35.0915 | -29.2584
19.2773 | -35.092 | -29.2654
19.2834 | -35.0925 | -29.2722
19.2895 | -35.093 | -29.2792
19.2956 | -35.0934 | -29.2864
19.3017 | -35.0939 | -29.2935
19.3078 | -35.0944 | -29.3005
19.3139 | -35.0949 | -29.3072
19.32 | -35.0954 | -29.3141
19.3261 | -35.0958 | -29.3211
19.3322 | -35.0963 | -29.3281
19.3383 | -35.0968 | -29.3349
19.3444 | -35.0973 | -29.3417
19.3505 | -35.0978 | -29.3484
19.3566 | -35.0982 | -29.3556
19.3627 | -35.0987 | -29.3626
19.3688 | -35.0991 | -29.3695
19.3749 | -35.0996 | -29.3766
19.381 | -35.1 | -29.3832
19.3871 | -35.1005 | -29.3902
19.3932 | -35.101 | -29.3973
19.3993 | -35.1014 | -29.4043
19.4054 | -35.1019 | -29.4114
19.4115 | -35.1024 | -29.4184
19.4177 | -35.1028 | -29.4252
19.4238 | -35.1033 | -29.432
19.4299 | -35.1038 | -29.439
19.436 | -35.1043 | -29.4466
19.4421 | -35.1048 | -29.4532
19.4482 | -35.1053 | -29.4601
19.4543 | -35.1058 | -29.4668
19.4604 | -35.1062 | -29.4735
19.4665 | -35.1067 | -29.4804
19.4726 | -35.1072 | -29.4871
19.4787 | -35.1077 | -29.494
19.4848 | -35.1081 | -29.5008
19.4909 | -35.1086 | -29.5074
19.497 | -35.1091 | -29.5141
19.5031 | -35.1094 | -29.521
19.5092 | -35.1099 | -29.5277
19.5153 | -35.1104 | -29.5346
19.5214 | -35.1109 | -29.5414
19.5275 | -35.1114 | -29.5483
19.5336 | -35.1118 | -29.555
19.5397 | -35.1123 | -29.5619
19.5458 | -35.1128 | -29.5688
19.5519 | -35.1133 | -29.5755
19.5581 | -35.1137 | -29.5819
19.5642 | -35.1142 | -29.5886
19.5703 | -35.1147 | -29.5955
19.5764 | -35.1152 | -29.6025
19.5825 | -35.1156 | -29.6093
19.5886 | -35.1161 | -29.6163
19.5947 | -35.1166 | -29.6228
19.6008 | -35.117 | -29.6296
19.6069 | -35.1175 | -29.6362
19.613 | -35.1179 | -29.6434
19.6191 | -35.1184 | -29.6505
19.6252 | -35.1189 | -29.6575
19.6313 | -35.1194 | -29.664
19.6374 | -35.1198 | -29.6711
19.6435 | -35.1203 | -29.6778
19.6496 | -35.1208 | -29.6843
19.6557 | -35.1213 | -29.6911
19.6618 | -35.1217 | -29.6977
19.6679 | -35.1221 | -29.7048
19.674 | -35.1226 | -29.7114
19.6801 | -35.123 | -29.7184
19.6862 | -35.1236 | -29.7251
19.6923 | -35.124 | -29.7319
19.6984 | -35.1246 | -29.7382
19.7046 | -35.1251 | -29.7453
19.7107 | -35.1256 | -29.752
19.7168 | -35.126 | -29.759
19.7229 | -35.1264 | -29.766
19.729 | -35.1269 | -29.7728
19.7351 | -35.1273 | -29.7793
19.7412 | -35.1278 | -29.7862
19.7473 | -35.1282 | -29.7931
19.7534 | -35.1287 | -29.8002
19.7595 | -35.1292 | -29.8071
19.7656 | -35.1297 | -29.8143
19.7717 | -35.1302 | -29.8207
19.7778 | -35.1306 | -29.8277
19.7839 | -35.1311 | -29.8344
19.79 | -35.1316 | -29.8415
19.7961 | -35.132 | -29.8487
19.8022 | -35.1324 | -29.8554
19.8083 | -35.133 | -29.8622
19.8144 | -35.1334 | -29.869
19.8205 | -35.1339 | -29.8757
19.8266 | -35.1344 | -29.8825
19.8327 | -35.1348 | -29.8893
19.8388 | -35.1353 | -29.8963
19.845 | -35.1358 | -29.9029
19.8511 | -35.1362 | -29.9101
19.8572 | -35.1367 | -29.9171
19.8633 | -35.1372 | -29.9239
19.8694 | -35.1377 | -29.9309
19.8755 | -35.1381 | -29.9377
19.8816 | -35.1387 | -29.9447
19.8877 | -35.1391 | -29.9516
19.8938 | -35.1396 | -29.9583
19.8999 | -35.14 | -29.965
19.906 | -35.1405 | -29.972
19.9121 | -35.141 | -29.9785
19.9182 | -35.1415 | -29.985
19.9243 | -35.142 | -29.9922
19.9304 | -35.1425 | -29.9994
19.9365 | -35.1429 | -30.0064
19.9426 | -35.1434 | -30.013
19.9487 | -35.1439 | -30.0195
19.9548 | -35.1444 | -30.0266
19.9609 | -35.1449 | -30.0333
19.967 | -35.1454 | -30.0397
19.9731 | -35.1459 | -30.0465
19.9792 | -35.1463 | -30.053
19.9853 | -35.1468 | -30.0598
19.9915 | -35.1473 | -30.0666
19.9976 | -35.1478 | -30.0729
20.0037 | -35.1483 | -30.0798
20.0098 | -35.1487 | -30.0864
20.0159 | -35.1491 | -30.0933
20.022 | -35.1496 | -30.1004
20.0281 | -35.1501 | -30.1073
20.0342 | -35.1506 | -30.1139
20.0403 | -35.151 | -30.1201
20.0464 | -35.1515 | -30.1273
20.0525 | -35.152 | -30.1337
20.0586 | -35.1525 | -30.1402
20.0647 | -35.153 | -30.1473
20.0708 | -35.1535 | -30.1546
20.0769 | -35.1539 | -30.1614
20.083 | -35.1544 | -30.1678
20.0891 | -35.1548 | -30.1747
20.0952 | -35.1553 | -30.1813
20.1013 | -35.1557 | -30.1878
20.1074 | -35.1562 | -30.1949
20.1135 | -35.1567 | -30.2013
20.1196 | -35.1572 | -30.2082
20.1257 | -35.1576 | -30.215
20.1319 | -35.1581 | -30.2217
20.138 | -35.1586 | -30.2283
20.1441 | -35.1591 | -30.2348
20.1502 | -35.1595 | -30.2417
20.1563 | -35.1601 | -30.2479
20.1624 | -35.1605 | -30.2545
20.1685 | -35.161 | -30.2613
20.1746 | -35.1614 | -30.2681
20.1807 | -35.1619 | -30.275
20.1868 | -35.1624 | -30.2817
20.1929 | -35.1628 | -30.2881
20.199 | -35.1632 | -30.2951
20.2051 | -35.1637 | -30.3016
20.2112 | -35.1642 | -30.3081
20.2173 | -35.1646 | -30.315
20.2234 | -35.165 | -30.3218
20.2295 | -35.1655 | -30.3278
20.2356 | -35.166 | -30.3345
20.2417 | -35.1665 | -30.3412
20.2478 | -35.1669 | -30.348
20.2539 | -35.1674 | -30.3553
20.26 | -35.1678 | -30.362
20.2661 | -35.1683 | -30.3689
20.2723 | -35.1688 | -30.3756
20.2784 | -35.1693 | -30.3824
20.2845 | -35.1697 | -30.3894
20.2906 | -35.1702 | -30.3962
20.2967 | -35.1707 | -30.4027
20.3028 | -35.1712 | -30.4095
20.3089 | -35.1717 | -30.4164
20.315 | -35.1721 | -30.4233
20.3211 | -35.1726 | -30.4302
20.3272 | -35.1731 | -30.4368
20.3333 | -35.1736 | -30.4432
20.3394 | -35.174 | -30.4496
20.3455 | -35.1745 | -30.4562
20.3516 | -35.1749 | -30.4626
20.3577 | -35.1753 | -30.4692
20.3638 | -35.1758 | -30.4759
20.3699 | -35.1763 | -30.4826
20.376 | -35.1767 | -30.4893
20.3821 | -35.1772 | -30.4955
20.3882 | -35.1776 | -30.5023
20.3943 | -35.1781 | -30.509
20.4004 | -35.1786 | -30.5153
20.4065 | -35.1791 | -30.522
20.4126 | -35.1795 | -30.5286
20.4188 | -35.18 | -30.5353
20.4249 | -35.1805 | -30.5418
20.431 | -35.1809 | -30.5487
20.4371 | -35.1814 | -30.5548
20.4432 | -35.1819 | -30.5616
20.4493 | -35.1823 | -30.5683
20.4554 | -35.1828 | -30.575
20.4615 | -35.1833 | -30.582
20.4676 | -35.1837 | -30.5886
20.4737 | -35.1842 | -30.5949
20.4798 | -35.1847 | -30.6015
20.4859 | -35.1852 | -30.6078
20.492 | -35.1856 | -30.6146
20.4981 | -35.1861 | -30.6211
20.5042 | -35.1865 | -30.6275
20.5103 | -35.1869 | -30.634
20.5164 | -35.1874 | -30.6407
20.5225 | -35.1878 | -30.6476
20.5286 | -35.1882 | -30.6545
20.5347 | -35.1886 | -30.6609
20.5408 | -35.1891 | -30.6679
20.5469 | -35.1896 | -30.6743
20.553 | -35.1901 | -30.6809
20.5592 | -35.1905 | -30.6875
20.5653 | -35.1909 | -30.6941
20.5714 | -35.1914 | -30.7005
20.5775 | -35.1919 | -30.7067
20.5836 | -35.1923 | -30.7135
20.5897 | -35.1928 | -30.7201
20.5958 | -35.1932 | -30.7267
20.6019 | -35.1936 | -30.7332
20.608 | -35.1941 | -30.7401
20.6141 | -35.1946 | -30.7465
20.6202 | -35.1951 | -30.7532
20.6263 | -35.1956 | -30.7595
20.6324 | -35.196 | -30.766
20.6385 | -35.1965 | -30.7723
20.6446 | -35.1969 | -30.779
20.6507 | -35.1973 | -30.7851
20.6568 | -35.1978 | -30.7919
20.6629 | -35.1983 | -30.7984
20.669 | -35.1988 | -30.8045
20.6751 | -35.1992 | -30.8111
20.6812 | -35.1996 | -30.8174
20.6873 | -35.2001 | -30.8242
20.6934 | -35.2005 | -30.8309
20.6995 | -35.201 | -30.8374
20.7057 | -35.2015 | -30.844
20.7118 | -35.202 | -30.8504
20.7179 | -35.2024 | -30.8564
20.724 | -35.2029 | -30.8628
20.7301 | -35.2034 | -30.8695
20.7362 | -35.2039 | -30.8763
20.7423 | -35.2043 | -30.8827
20.7484 | -35.2048 | -30.8894
20.7545 | -35.2053 | -30.8957
20.7606 | -35.2058 | -30.9022
20.7667 | -35.2062 | -30.9089
20.7728 | -35.2067 | -30.9155
20.7789 | -35.2072 | -30.9223
20.785 | -35.2077 | -30.9287
20.7911 | -35.2082 | -30.9354
20.7972 | -35.2086 | -30.9419
20.8033 | -35.2091 | -30.9485
20.8094 | -35.2095 | -30.9554
20.8155 | -35.2099 | -30.9618
20.8216 | -35.2104 | -30.9686
20.8277 | -35.2109 | -30.9748
20.8338 | -35.2114 | -30.9816
20.8399 | -35.2118 | -30.9881
20.8461 | -35.2123 | -30.9945
20.8522 | -35.2128 | -31.001
20.8583 | -35.2133 | -31.0073
20.8644 | -35.2137 | -31.0138
20.8705 | -35.2142 | -31.0209
20.8766 | -35.2147 | -31.028
20.8827 | -35.2152 | -31.0345
20.8888 | -35.2156 | -31.0408
20.8949 | -35.2162 | -31.0476
20.901 | -35.2166 | -31.054
20.9071 | -35.217 | -31.0604
20.9132 | -35.2175 | -31.067
20.9193 | -35.218 | -31.0734
20.9254 | -35.2185 | -31.0799
20.9315 | -35.219 | -31.0862
20.9376 | -35.2194 | -31.0929
20.9437 | -35.2198 | -31.0995
20.9498 | -35.2202 | -31.106
20.9559 | -35.2207 | -31.1122
20.962 | -35.2213 | -31.119
20.9681 | -35.2217 | -31.1255
20.9742 | -35.2221 | -31.132
20.9803 | -35.2226 | -31.1387
20.9864 | -35.223 | -31.1453
20.9926 | -35.2235 | -31.152
20.9987 | -35.2239 | -31.1583
21.0048 | -35.2243 | -31.1649
21.0109 | -35.2247 | -31.1713
21.017 | -35.2252 | -31.1775
21.0231 | -35.2256 | -31.1841
21.0292 | -35.2261 | -31.1909
21.0353 | -35.2266 | -31.1966
21.0414 | -35.227 | -31.203
21.0475 | -35.2275 | -31.2094
21.0536 | -35.228 | -31.2157
21.0597 | -35.2285 | -31.2218
21.0658 | -35.2289 | -31.228
21.0719 | -35.2293 | -31.2346
21.078 | -35.2298 | -31.2416
21.0841 | -35.2304 | -31.2481
21.0902 | -35.231 | -31.2546
21.0963 | -35.2314 | -31.2611
21.1024 | -35.2318 | -31.2673
21.1085 | -35.2321 | -31.2739
21.1146 | -35.2326 | -31.2805
21.1207 | -35.233 | -31.2868
21.1268 | -35.2335 | -31.2933
21.133 | -35.234 | -31.2998
21.1391 | -35.2344 | -31.3056
21.1452 | -35.2349 | -31.3123
21.1513 | -35.2354 | -31.3187
21.1574 | -35.2359 | -31.3255
21.1635 | -35.2363 | -31.3317
21.1696 | -35.2368 | -31.3385
21.1757 | -35.2372 | -31.345
21.1818 | -35.2377 | -31.3516
21.1879 | -35.2381 | -31.3582
21.194 | -35.2386 | -31.3646
21.2001 | -35.239 | -31.3709
21.2062 | -35.2394 | -31.3774
21.2123 | -35.2398 | -31.3837
21.2184 | -35.2403 | -31.39
21.2245 | -35.2407 | -31.3964
21.2306 | -35.2412 | -31.4026
21.2367 | -35.2416 | -31.4091
21.2428 | -35.2421 | -31.4155
21.2489 | -35.2425 | -31.4215
21.255 | -35.2429 | -31.428
21.2611 | -35.2434 | -31.4341
21.2672 | -35.2438 | -31.441
21.2733 | -35.2442 | -31.4473
21.2795 | -35.2447 | -31.4538
21.2856 | -35.2452 | -31.4601
21.2917 | -35.2456 | -31.4671
21.2978 | -35.246 | -31.4741
21.3039 | -35.2465 | -31.4804
21.31 | -35.2469 | -31.4868
21.3161 | -35.2474 | -31.4932
21.3222 | -35.2479 | -31.4994
21.3283 | -35.2483 | -31.5061
21.3344 | -35.2487 | -31.5118
21.3405 | -35.2492 | -31.5184
21.3466 | -35.2497 | -31.5246
21.3527 | -35.2502 | -31.5306
21.3588 | -35.2507 | -31.537
21.3649 | -35.2511 | -31.5434
21.371 | -35.2516 | -31.5497
21.3771 | -35.252 | -31.5559
21.3832 | -35.2526 | -31.5619
21.3893 | -35.253 | -31.5678
21.3954 | -35.2534 | -31.5741
21.4015 | -35.2539 | -31.5803
21.4076 | -35.2543 | -31.5867
21.4137 | -35.2547 | -31.593
21.4199 | -35.2552 | -31.5994
21.426 | -35.2555 | -31.6061
21.4321 | -35.256 | -31.6122
21.4382 | -35.2564 | -31.618
21.4443 | -35.2569 | -31.6245
21.4504 | -35.2574 | -31.6316
21.4565 | -35.2578 | -31.6379
21.4626 | -35.2583 | -31.6443
21.4687 | -35.2588 | -31.6506
21.4748 | -35.2592 | -31.6565
21.4809 | -35.2596 | -31.6627
21.487 | -35.26 | -31.6694
21.4931 | -35.2605 | -31.6753
21.4992 | -35.2609 | -31.6814
21.5053 | -35.2613 | -31.6874
21.5114 | -35.2618 | -31.6938
21.5175 | -35.2622 | -31.7004
21.5236 | -35.2626 | -31.7064
21.5297 | -35.263 | -31.7124
21.5358 | -35.2635 | -31.719
21.5419 | -35.2639 | -31.7257
21.548 | -35.2644 | -31.732
21.5541 | -35.2648 | -31.7381
21.5602 | -35.2652 | -31.7446
21.5664 | -35.2656 | -31.7498
21.5725 | -35.266 | -31.7566
21.5786 | -35.2665 | -31.7633
21.5847 | -35.2669 | -31.7696
21.5908 | -35.2674 | -31.7757
21.5969 | -35.2678 | -31.782
21.603 | -35.2682 | -31.7883
21.6091 | -35.2687 | -31.7948
21.6152 | -35.2692 | -31.8017
21.6213 | -35.2696 | -31.8087
21.6274 | -35.2701 | -31.8155
21.6335 | -35.2705 | -31.8223
21.6396 | -35.271 | -31.8287
21.6457 | -35.2714 | -31.8355
21.6518 | -35.2717 | -31.842
21.6579 | -35.2722 | -31.848
21.664 | -35.2726 | -31.8543
21.6701 | -35.273 | -31.8598
21.6762 | -35.2734 | -31.8666
21.6823 | -35.2738 | -31.8729
21.6884 | -35.2741 | -31.8796
21.6945 | -35.2746 | -31.8856
21.7006 | -35.275 | -31.8925
21.7068 | -35.2754 | -31.8989
21.7129 | -35.2758 | -31.9053
21.719 | -35.2763 | -31.9115
21.7251 | -35.2767 | -31.9179
21.7312 | -35.2772 | -31.9244
21.7373 | -35.2777 | -31.9311
21.7434 | -35.2781 | -31.9371
21.7495 | -35.2786 | -31.9435
21.7556 | -35.2791 | -31.9499
21.7617 | -35.2796 | -31.9564
21.7678 | -35.28 | -31.9629
21.7739 | -35.2805 | -31.9698
21.78 | -35.2809 | -31.9761
21.7861 | -35.2813 | -31.9826
21.7922 | -35.2818 | -31.9889
21.7983 | -35.2822 | -31.995
21.8044 | -35.2827 | -32.0014
21.8105 | -35.2832 | -32.007
21.8166 | -35.2837 | -32.0136
21.8227 | -35.2842 | -32.02
21.8288 | -35.2847 | -32.0263
21.8349 | -35.2851 | -32.0324
21.841 | -35.2856 | -32.0391
21.8471 | -35.286 | -32.046
21.8533 | -35.2863 | -32.0521
21.8594 | -35.2868 | -32.0584
21.8655 | -35.2873 | -32.0651
21.8716 | -35.2877 | -32.0719
21.8777 | -35.2882 | -32.078
21.8838 | -35.2887 | -32.0843
21.8899 | -35.2891 | -32.0904
21.896 | -35.2895 | -32.097
21.9021 | -35.29 | -32.1034
21.9082 | -35.2904 | -32.1102
21.9143 | -35.2909 | -32.1168
21.9204 | -35.2913 | -32.1228
21.9265 | -35.2917 | -32.1291
21.9326 | -35.2921 | -32.1354
21.9387 | -35.2925 | -32.1418
21.9448 | -35.293 | -32.1483
21.9509 | -35.2934 | -32.1546
21.957 | -35.2939 | -32.1608
21.9631 | -35.2944 | -32.1674
21.9692 | -35.2949 | -32.1736
21.9753 | -35.2954 | -32.1799
21.9814 | -35.2958 | -32.1862
21.9875 | -35.2962 | -32.1924
21.9937 | -35.2966 | -32.1985
21.9998 | -35.2971 | -32.2049
22.0059 | -35.2975 | -32.2116
22.012 | -35.298 | -32.2178
22.0181 | -35.2985 | -32.2237
22.0242 | -35.299 | -32.2298
22.0303 | -35.2994 | -32.2361
22.0364 | -35.2999 | -32.2421
22.0425 | -35.3003 | -32.2483
22.0486 | -35.3007 | -32.2546
22.0547 | -35.3012 | -32.2603
22.0608 | -35.3016 | -32.2663
22.0669 | -35.302 | -32.2722
22.073 | -35.3025 | -32.2779
22.0791 | -35.3029 | -32.2842
22.0852 | -35.3033 | -32.2907
22.0913 | -35.3037 | -32.297
22.0974 | -35.3041 | -32.3037
22.1035 | -35.3045 | -32.3096
22.1096 | -35.3049 | -32.3161
22.1157 | -35.3053 | -32.322
22.1218 | -35.3058 | -32.3281
22.1279 | -35.3061 | -32.3346
22.134 | -35.3066 | -32.3406
22.1402 | -35.3071 | -32.3468
22.1463 | -35.3075 | -32.3534
22.1524 | -35.3079 | -32.3597
22.1585 | -35.3083 | -32.3663
22.1646 | -35.3087 | -32.3728
22.1707 | -35.3091 | -32.3792
22.1768 | -35.3096 | -32.386
22.1829 | -35.3101 | -32.3924
22.189 | -35.3105 | -32.3989
22.1951 | -35.311 | -32.4053
22.2012 | -35.3115 | -32.4117
22.2073 | -35.3119 | -32.4181
22.2134 | -35.3123 | -32.4246
22.2195 | -35.3127 | -32.4306
22.2256 | -35.3131 | -32.4367
22.2317 | -35.3136 | -32.4427
22.2378 | -35.314 | -32.4488
22.2439 | -35.3144 | -32.4557
22.25 | -35.3149 | -32.4621
22.2561 | -35.3153 | -32.4688
22.2622 | -35.3157 | -32.4751
22.2683 | -35.3161 | -32.4811
22.2744 | -35.3165 | -32.4876
22.2806 | -35.317 | -32.4937
22.2867 | -35.3174 | -32.5003
22.2928 | -35.3178 | -32.5062
22.2989 | -35.3182 | -32.5125
22.305 | -35.3186 | -32.5184
22.3111 | -35.319 | -32.5249
22.3172 | -35.3195 | -32.5311
22.3233 | -35.3199 | -32.5372
22.3294 | -35.3204 | -32.5435
22.3355 | -35.3208 | -32.5498
22.3416 | -35.3213 | -32.5563
22.3477 | -35.3218 | -32.5625
22.3538 | -35.3222 | -32.5685
22.3599 | -35.3226 | -32.5746
22.366 | -35.323 | -32.5808
22.3721 | -35.3235 | -32.5866
22.3782 | -35.3239 | -32.5926
22.3843 | -35.3244 | -32.5992
22.3904 | -35.3248 | -32.6057
22.3965 | -35.3252 | -32.6118
22.4026 | -35.3257 | -32.6179
22.4087 | -35.3261 | -32.6241
22.4148 | -35.3265 | -32.6302
22.4209 | -35.3269 | -32.6365
22.4271 | -35.3273 | -32.6429
22.4332 | -35.3277 | -32.6487
22.4393 | -35.3281 | -32.6551
22.4454 | -35.3285 | -32.6613
22.4515 | -35.3289 | -32.6675
22.4576 | -35.3293 | -32.6743
22.4637 | -35.3297 | -32.6809
22.4698 | -35.3302 | -32.6873
22.4759 | -35.3306 | -32.6938
22.482 | -35.331 | -32.7002
22.4881 | -35.3314 | -32.7064
22.4942 | -35.3319 | -32.7127
22.5003 | -35.3323 | -32.7191
22.5064 | -35.3327 | -32.7257
22.5125 | -35.3331 | -32.7315
22.5186 | -35.3335 | -32.7377
22.5247 | -35.3339 | -32.7442
22.5308 | -35.3343 | -32.7504
22.5369 | -35.3348 | -32.757
22.543 | -35.3352 | -32.7632
22.5491 | -35.3357 | -32.7696
22.5552 | -35.3361 | -32.7761
22.5613 | -35.3366 | -32.7824
22.5675 | -35.337 | -32.7886
22.5736 | -35.3374 | -32.7951
22.5797 | -35.3379 | -32.8011
22.5858 | -35.3384 | -32.8074
22.5919 | -35.3389 | -32.8135
22.598 | -35.3393 | -32.8195
22.6041 | -35.3397 | -32.8256
22.6102 | -35.3402 | -32.832
22.6163 | -35.3405 | -32.8386
22.6224 | -35.341 | -32.8452
22.6285 | -35.3414 | -32.8512
22.6346 | -35.3418 | -32.8574
22.6407 | -35.3422 | -32.8643
22.6468 | -35.3426 | -32.8704
22.6529 | -35.343 | -32.876
22.659 | -35.3434 | -32.8825
22.6651 | -35.3439 | -32.889
22.6712 | -35.3443 | -32.8951
22.6773 | -35.3447 | -32.9012
22.6834 | -35.3452 | -32.9073
22.6895 | -35.3457 | -32.9132
22.6956 | -35.3462 | -32.9195
22.7017 | -35.3466 | -32.9255
22.7079 | -35.347 | -32.932
22.714 | -35.3474 | -32.9384
22.7201 | -35.3478 | -32.9444
22.7262 | -35.3482 | -32.9508
22.7323 | -35.3487 | -32.9569
22.7384 | -35.3492 | -32.9631
22.7445 | -35.3496 | -32.9692
22.7506 | -35.3499 | -32.9754
22.7567 | -35.3504 | -32.9814
22.7628 | -35.3508 | -32.9877
22.7689 | -35.3512 | -32.9936
22.775 | -35.3515 | -33.0001
22.7811 | -35.352 | -33.0065
22.7872 | -35.3524 | -33.0129
22.7933 | -35.3528 | -33.0189
22.7994 | -35.3533 | -33.0254
22.8055 | -35.3537 | -33.0321
22.8116 | -35.3541 | -33.0381
22.8177 | -35.3545 | -33.0444
22.8238 | -35.3549 | -33.0506
22.8299 | -35.3553 | -33.0565
22.836 | -35.3557 | -33.0628
22.8421 | -35.3562 | -33.069
22.8482 | -35.3566 | -33.0752
22.8544 | -35.357 | -33.0818
22.8605 | -35.3574 | -33.0879
22.8666 | -35.3578 | -33.0941
22.8727 | -35.3582 | -33.1002
22.8788 | -35.3587 | -33.1066
22.8849 | -35.3591 | -33.1126
22.891 | -35.3595 | -33.119
22.8971 | -35.3599 | -33.1246
22.9032 | -35.3605 | -33.1313
22.9093 | -35.3609 | -33.1375
22.9154 | -35.3614 | -33.1436
22.9215 | -35.3618 | -33.1499
22.9276 | -35.3622 | -33.1562
22.9337 | -35.3627 | -33.1621
22.9398 | -35.3631 | -33.1681
22.9459 | -35.3635 | -33.1751
22.952 | -35.3639 | -33.1815
22.9581 | -35.3643 | -33.188
22.9642 | -35.3647 | -33.1943
22.9703 | -35.3651 | -33.2009
22.9764 | -35.3655 | -33.2074
22.9825 | -35.366 | -33.2133
22.9886 | -35.3664 | -33.22
22.9948 | -35.3667 | -33.2261
23.0009 | -35.3672 | -33.2325
23.007 | -35.3676 | -33.2392
23.0131 | -35.3681 | -33.2456
23.0192 | -35.3685 | -33.2516
23.0253 | -35.369 | -33.2576
23.0314 | -35.3693 | -33.2638
23.0375 | -35.3698 | -33.2702
23.0436 | -35.3702 | -33.277
23.0497 | -35.3706 | -33.2831
23.0558 | -35.371 | -33.2891
23.0619 | -35.3714 | -33.2955
23.068 | -35.3719 | -33.3016
23.0741 | -35.3723 | -33.3074
23.0802 | -35.3727 | -33.3139
23.0863 | -35.3731 | -33.3199
23.0924 | -35.3735 | -33.3265
23.0985 | -35.3738 | -33.3327
23.1046 | -35.3742 | -33.3392
23.1107 | -35.3746 | -33.3454
23.1168 | -35.375 | -33.3516
23.1229 | -35.3754 | -33.3577
23.129 | -35.3759 | -33.364
23.1351 | -35.3763 | -33.3701
23.1413 | -35.3767 | -33.3763
23.1474 | -35.3771 | -33.3826
23.1535 | -35.3775 | -33.3885
23.1596 | -35.3779 | -33.3942
23.1657 | -35.3783 | -33.4005
23.1718 | -35.3788 | -33.4066
23.1779 | -35.3792 | -33.413
23.184 | -35.3796 | -33.4191
23.1901 | -35.38 | -33.425
23.1962 | -35.3805 | -33.4314
23.2023 | -35.3808 | -33.4374
23.2084 | -35.3812 | -33.4437
23.2145 | -35.3817 | -33.4496
23.2206 | -35.382 | -33.4556
23.2267 | -35.3824 | -33.4622
23.2328 | -35.3828 | -33.4685
23.2389 | -35.3832 | -33.4747
23.245 | -35.3836 | -33.4808
23.2511 | -35.3839 | -33.4862
23.2572 | -35.3844 | -33.4925
23.2633 | -35.3848 | -33.4988
23.2694 | -35.3854 | -33.5046
23.2755 | -35.3857 | -33.511
23.2817 | -35.3862 | -33.5174
23.2878 | -35.3865 | -33.5239
23.2939 | -35.3869 | -33.5302
23.3 | -35.3873 | -33.5366
23.3061 | -35.3877 | -33.5428
23.3122 | -35.3881 | -33.5496
23.3183 | -35.3884 | -33.556
23.3244 | -35.3888 | -33.5622
23.3305 | -35.3892 | -33.5676
23.3366 | -35.3896 | -33.5743
23.3427 | -35.39 | -33.5806
23.3488 | -35.3904 | -33.587
23.3549 | -35.3908 | -33.5932
23.361 | -35.3913 | -33.5994
23.3671 | -35.3916 | -33.606
23.3732 | -35.392 | -33.6116
23.3793 | -35.3924 | -33.6175
23.3854 | -35.3928 | -33.624
23.3915 | -35.3932 | -33.6304
23.3976 | -35.3936 | -33.6371
23.4037 | -35.394 | -33.643
23.4098 | -35.3944 | -33.6489
23.4159 | -35.3948 | -33.6556
23.422 | -35.3952 | -33.6618
23.4282 | -35.3956 | -33.6674
23.4343 | -35.396 | -33.6735
23.4404 | -35.3964 | -33.6802
23.4465 | -35.3967 | -33.686
23.4526 | -35.3971 | -33.692
23.4587 | -35.3975 | -33.6981
23.4648 | -35.3978 | -33.7044
23.4709 | -35.3983 | -33.7106
23.477 | -35.3986 | -33.717
23.4831 | -35.3991 | -33.7231
23.4892 | -35.3996 | -33.7295
23.4953 | -35.4 | -33.7358
23.5014 | -35.4003 | -33.7421
23.5075 | -35.4006 | -33.7483
23.5136 | -35.401 | -33.7548
23.5197 | -35.4014 | -33.761
23.5258 | -35.4019 | -33.7672
23.5319 | -35.4023 | -33.7736
23.538 | -35.4027 | -33.7798
23.5441 | -35.4031 | -33.7863
23.5502 | -35.4034 | -33.7927
23.5563 | -35.4038 | -33.7987
23.5624 | -35.4042 | -33.8051
23.5686 | -35.4046 | -33.811
23.5747 | -35.4049 | -33.8168
23.5808 | -35.4054 | -33.8229
23.5869 | -35.4059 | -33.8292
23.593 | -35.4063 | -33.8353
23.5991 | -35.4066 | -33.8414
23.6052 | -35.407 | -33.8474
23.6113 | -35.4074 | -33.8535
23.6174 | -35.4077 | -33.8599
23.6235 | -35.4081 | -33.8662
23.6296 | -35.4085 | -33.8727
23.6357 | -35.4089 | -33.8788
23.6418 | -35.4093 | -33.8845
23.6479 | -35.4098 | -33.8909
23.654 | -35.4102 | -33.8969
23.6601 | -35.4105 | -33.9029
23.6662 | -35.411 | -33.9092
23.6723 | -35.4114 | -33.9156
23.6784 | -35.4118 | -33.9216
23.6845 | -35.4122 | -33.928
23.6906 | -35.4126 | -33.9341
23.6967 | -35.413 | -33.9401
23.7028 | -35.4134 | -33.9464
23.7089 | -35.4138 | -33.9524
23.7151 | -35.4143 | -33.9586
23.7212 | -35.4147 | -33.9645
23.7273 | -35.415 | -33.9707
23.7334 | -35.4155 | -33.9771
23.7395 | -35.4159 | -33.9828
23.7456 | -35.4162 | -33.9884
23.7517 | -35.4167 | -33.9943
23.7578 | -35.4171 | -34.0004
23.7639 | -35.4174 | -34.0062
23.77 | -35.4179 | -34.0121
23.7761 | -35.4182 | -34.0183
23.7822 | -35.4186 | -34.0243
23.7883 | -35.4191 | -34.0301
23.7944 | -35.4194 | -34.0365
23.8005 | -35.4198 | -34.0435
23.8066 | -35.4202 | -34.0502
23.8127 | -35.4206 | -34.0566
23.8188 | -35.421 | -34.0632
23.8249 | -35.4213 | -34.0691
23.831 | -35.4217 | -34.0754
23.8371 | -35.4221 | -34.081
23.8432 | -35.4225 | -34.0864
23.8493 | -35.4229 | -34.0922
23.8555 | -35.4233 | -34.0981
23.8616 | -35.4237 | -34.1036
23.8677 | -35.4242 | -34.1101
23.8738 | -35.4246 | -34.1159
23.8799 | -35.4249 | -34.1221
23.886 | -35.4253 | -34.1279
23.8921 | -35.4256 | -34.1341
23.8982 | -35.426 | -34.1405
23.9043 | -35.4264 | -34.1465
23.9104 | -35.4268 | -34.153
23.9165 | -35.4272 | -34.159
23.9226 | -35.4275 | -34.1651
23.9287 | -35.4279 | -34.1714
23.9348 | -35.4283 | -34.1778
23.9409 | -35.4287 | -34.1842
23.947 | -35.4291 | -34.1901
23.9531 | -35.4295 | -34.1957
23.9592 | -35.4299 | -34.2022
23.9653 | -35.4302 | -34.2087
23.9714 | -35.4305 | -34.2147
23.9775 | -35.4309 | -34.221
23.9836 | -35.4313 | -34.2275
23.9897 | -35.4317 | -34.2337
23.9958 | -35.432 | -34.2396
24.002 | -35.4324 | -34.2458
24.0081 | -35.4328 | -34.2519
24.0142 | -35.4332 | -34.2584
24.0203 | -35.4336 | -34.2644
24.0264 | -35.4341 | -34.2705
24.0325 | -35.4344 | -34.2763
24.0386 | -35.4347 | -34.2827
24.0447 | -35.4351 | -34.289
24.0508 | -35.4356 | -34.2955
24.0569 | -35.436 | -34.3015
24.063 | -35.4365 | -34.3077
24.0691 | -35.4369 | -34.3135
24.0752 | -35.4372 | -34.3197
24.0813 | -35.4376 | -34.3261
24.0874 | -35.4381 | -34.3317
24.0935 | -35.4385 | -34.3381
24.0996 | -35.4388 | -34.3442
24.1057 | -35.4392 | -34.3498
24.1118 | -35.4396 | -34.3561
24.1179 | -35.44 | -34.3622
24.124 | -35.4403 | -34.368
24.1301 | -35.4407 | -34.3737
24.1362 | -35.4411 | -34.3802
24.1424 | -35.4414 | -34.3861
24.1485 | -35.4418 | -34.3921
24.1546 | -35.4422 | -34.3982
24.1607 | -35.4425 | -34.404
24.1668 | -35.4429 | -34.4102
24.1729 | -35.4433 | -34.416
24.179 | -35.4436 | -34.422
24.1851 | -35.444 | -34.4284
24.1912 | -35.4444 | -34.4342
24.1973 | -35.4448 | -34.4404
24.2034 | -35.4452 | -34.4464
24.2095 | -35.4455 | -34.4529
24.2156 | -35.446 | -34.4592
24.2217 | -35.4464 | -34.4654
24.2278 | -35.4468 | -34.4714
24.2339 | -35.4472 | -34.4773
24.24 | -35.4475 | -34.4831
24.2461 | -35.4479 | -34.4892
24.2522 | -35.4482 | -34.4951
24.2583 | -35.4487 | -34.501
24.2644 | -35.449 | -34.5075
24.2705 | -35.4494 | -34.5137
24.2766 | -35.4497 | -34.52
24.2827 | -35.4501 | -34.5257
24.2889 | -35.4505 | -34.5316
24.295 | -35.4509 | -34.5376
24.3011 | -35.4513 | -34.5438
24.3072 | -35.4516 | -34.5499
24.3133 | -35.452 | -34.5556
24.3194 | -35.4523 | -34.5619
24.3255 | -35.4527 | -34.568
24.3316 | -35.4532 | -34.5741
24.3377 | -35.4535 | -34.58
24.3438 | -35.4538 | -34.5856
24.3499 | -35.4542 | -34.5918
24.356 | -35.4545 | -34.598
24.3621 | -35.4549 | -34.6038
24.3682 | -35.4553 | -34.6098
24.3743 | -35.4557 | -34.6159
24.3804 | -35.456 | -34.6213
24.3865 | -35.4564 | -34.6273
24.3926 | -35.4567 | -34.6328
24.3987 | -35.4571 | -34.6395
24.4048 | -35.4575 | -34.6457
24.4109 | -35.4578 | -34.6513
24.417 | -35.4582 | -34.6568
24.4231 | -35.4585 | -34.663
24.4293 | -35.459 | -34.6689
24.4354 | -35.4593 | -34.6747
24.4415 | -35.4597 | -34.6803
24.4476 | -35.4601 | -34.6859
24.4537 | -35.4604 | -34.6916
24.4598 | -35.4608 | -34.6981
24.4659 | -35.4612 | -34.704
24.472 | -35.4616 | -34.71
24.4781 | -35.462 | -34.716
24.4842 | -35.4623 | -34.7216
24.4903 | -35.4626 | -34.7275
24.4964 | -35.4629 | -34.7334
24.5025 | -35.4633 | -34.7393
24.5086 | -35.4637 | -34.7449
24.5147 | -35.4641 | -34.7506
24.5208 | -35.4645 | -34.7563
24.5269 | -35.4648 | -34.7621
24.533 | -35.4652 | -34.7683
24.5391 | -35.4655 | -34.7746
24.5452 | -35.4659 | -34.7802
24.5513 | -35.4663 | -34.7858
24.5574 | -35.4666 | -34.7919
24.5635 | -35.467 | -34.7975
24.5696 | -35.4673 | -34.8034
24.5758 | -35.4677 | -34.8095
24.5819 | -35.468 | -34.8151
24.588 | -35.4683 | -34.8211
24.5941 | -35.4687 | -34.8269
24.6002 | -35.4691 | -34.833
24.6063 | -35.4694 | -34.8388
24.6124 | -35.4699 | -34.845
24.6185 | -35.4703 | -34.8508
24.6246 | -35.4706 | -34.8566
24.6307 | -35.471 | -34.8625
24.6368 | -35.4713 | -34.8684
24.6429 | -35.4717 | -34.8741
24.649 | -35.4721 | -34.88
24.6551 | -35.4724 | -34.8856
24.6612 | -35.4727 | -34.8911
24.6673 | -35.4731 | -34.897
24.6734 | -35.4736 | -34.9031
24.6795 | -35.474 | -34.9091
24.6856 | -35.4744 | -34.9153
24.6917 | -35.4747 | -34.9216
24.6978 | -35.475 | -34.928
24.7039 | -35.4753 | -34.9337
24.71 | -35.4757 | -34.9396
24.7162 | -35.4761 | -34.9455
24.7223 | -35.4765 | -34.9514
24.7284 | -35.4769 | -34.9571
24.7345 | -35.4772 | -34.9629
24.7406 | -35.4776 | -34.9689
24.7467 | -35.478 | -34.975
24.7528 | -35.4783 | -34.9803
24.7589 | -35.4786 | -34.9856
24.765 | -35.4789 | -34.9909
24.7711 | -35.4792 | -34.9967
24.7772 | -35.4795 | -35.0027
24.7833 | -35.4799 | -35.0086
24.7894 | -35.4803 | -35.0144
24.7955 | -35.4806 | -35.0206
24.8016 | -35.481 | -35.0263
24.8077 | -35.4813 | -35.0321
24.8138 | -35.4818 | -35.0381
24.8199 | -35.482 | -35.0443
24.826 | -35.4824 | -35.0498
24.8321 | -35.4828 | -35.0558
24.8382 | -35.4831 | -35.0614
24.8443 | -35.4835 | -35.0672
24.8504 | -35.4839 | -35.0731
24.8565 | -35.4843 | -35.079
24.8627 | -35.4846 | -35.0852
24.8688 | -35.4849 | -35.0908
24.8749 | -35.4853 | -35.0964
24.881 | -35.4856 | -35.1023
24.8871 | -35.486 | -35.1079
24.8932 | -35.4864 | -35.1135
24.8993 | -35.4868 | -35.1191
24.9054 | -35.4871 | -35.1249
24.9115 | -35.4875 | -35.1305
24.9176 | -35.4878 | -35.1361
24.9237 | -35.4883 | -35.1421
24.9298 | -35.4886 | -35.1481
24.9359 | -35.4889 | -35.154
24.942 | -35.4893 | -35.1595
24.9481 | -35.4897 | -35.1654
24.9542 | -35.49 | -35.1711
24.9603 | -35.4904 | -35.1771
24.9664 | -35.4907 | -35.1827
24.9725 | -35.4911 | -35.1886
24.9786 | -35.4913 | -35.1939
24.9847 | -35.4916 | -35.1998
24.9908 | -35.492 | -35.2057
24.9969 | -35.4924 | -35.2116
25.0031 | -35.4927 | -35.2178
25.0092 | -35.493 | -35.2239
25.0153 | -35.4934 | -35.2298
25.0214 | -35.4938 | -35.2359
25.0275 | -35.4941 | -35.2415
25.0336 | -35.4945 | -35.2474
25.0397 | -35.4948 | -35.2535
25.0458 | -35.4952 | -35.2593
25.0519 | -35.4955 | -35.2652
25.058 | -35.4959 | -35.2714
25.0641 | -35.4963 | -35.277
25.0702 | -35.4966 | -35.2826
25.0763 | -35.497 | -35.2883
25.0824 | -35.4973 | -35.294
25.0885 | -35.4977 | -35.2995
25.0946 | -35.4981 | -35.3052
25.1007 | -35.4984 | -35.311
25.1068 | -35.4988 | -35.3167
25.1129 | -35.4991 | -35.3224
25.119 | -35.4995 | -35.3277
25.1251 | -35.4998 | -35.333
25.1312 | -35.5001 | -35.3388
25.1373 | -35.5005 | -35.3448
25.1435 | -35.5009 | -35.3505
25.1496 | -35.5013 | -35.3558
25.1557 | -35.5016 | -35.3615
25.1618 | -35.502 | -35.3674
25.1679 | -35.5023 | -35.373
25.174 | -35.5027 | -35.3788
25.1801 | -35.503 | -35.3845
25.1862 | -35.5034 | -35.3901
25.1923 | -35.5037 | -35.3958
25.1984 | -35.5041 | -35.4018
25.2045 | -35.5044 | -35.4074
25.2106 | -35.5048 | -35.4132
25.2167 | -35.5051 | -35.4191
25.2228 | -35.5055 | -35.425
25.2289 | -35.5059 | -35.4306
25.235 | -35.5062 | -35.4364
25.2411 | -35.5065 | -35.4419
25.2472 | -35.5068 | -35.4475
25.2533 | -35.5072 | -35.4534
25.2594 | -35.5075 | -35.4588
25.2655 | -35.5078 | -35.4647
25.2716 | -35.5082 | -35.4701
25.2777 | -35.5086 | -35.4759
25.2838 | -35.5089 | -35.4819
25.29 | -35.5092 | -35.4876
25.2961 | -35.5095 | -35.4928
25.3022 | -35.5099 | -35.4984
25.3083 | -35.5102 | -35.504
25.3144 | -35.5105 | -35.5097
25.3205 | -35.5109 | -35.5151
25.3266 | -35.5112 | -35.5203
25.3327 | -35.5115 | -35.5262
25.3388 | -35.5119 | -35.5319
25.3449 | -35.5122 | -35.5374
25.351 | -35.5127 | -35.5432
25.3571 | -35.513 | -35.5485
25.3632 | -35.5134 | -35.5541
25.3693 | -35.5137 | -35.5594
25.3754 | -35.514 | -35.5651
25.3815 | -35.5144 | -35.5707
25.3876 | -35.5147 | -35.5766
25.3937 | -35.5151 | -35.582
25.3998 | -35.5154 | -35.5875
25.4059 | -35.5157 | -35.5931
25.412 | -35.516 | -35.5986
25.4181 | -35.5164 | -35.6046
25.4242 | -35.5167 | -35.6102
25.4304 | -35.517 | -35.6157
25.4365 | -35.5174 | -35.6213
25.4426 | -35.5177 | -35.6271
25.4487 | -35.518 | -35.6329
25.4548 | -35.5184 | -35.6385
25.4609 | -35.5187 | -35.6442
25.467 | -35.519 | -35.6498
25.4731 | -35.5194 | -35.6555
25.4792 | -35.5197 | -35.6611
25.4853 | -35.5201 | -35.6666
25.4914 | -35.5204 | -35.6722
25.4975 | -35.5208 | -35.6777
25.5036 | -35.5212 | -35.683
25.5097 | -35.5215 | -35.6882
25.5158 | -35.5218 | -35.6936
25.5219 | -35.5222 | -35.699
25.528 | -35.5225 | -35.7046
25.5341 | -35.5229 | -35.7102
25.5402 | -35.5233 | -35.7154
25.5463 | -35.5236 | -35.7212
25.5524 | -35.524 | -35.7271
25.5585 | -35.5243 | -35.7324
25.5646 | -35.5247 | -35.738
25.5707 | -35.525 | -35.7435
25.5769 | -35.5253 | -35.7492
25.583 | -35.5256 | -35.755
25.5891 | -35.526 | -35.7605
25.5952 | -35.5264 | -35.7662
25.6013 | -35.5268 | -35.7719
25.6074 | -35.5271 | -35.7771
25.6135 | -35.5274 | -35.7825
25.6196 | -35.5278 | -35.7885
25.6257 | -35.5281 | -35.7939
25.6318 | -35.5284 | -35.7992
25.6379 | -35.5288 | -35.8047
25.644 | -35.5291 | -35.8099
25.6501 | -35.5294 | -35.8153
25.6562 | -35.5298 | -35.8208
25.6623 | -35.5301 | -35.8264
25.6684 | -35.5304 | -35.832
25.6745 | -35.5307 | -35.8378
25.6806 | -35.5311 | -35.8427
25.6867 | -35.5314 | -35.8484
25.6928 | -35.5318 | -35.8544
25.6989 | -35.5321 | -35.8598
25.705 | -35.5324 | -35.8659
25.7111 | -35.5328 | -35.8714
25.7173 | -35.5331 | -35.8765
25.7234 | -35.5335 | -35.8822
25.7295 | -35.5338 | -35.8877
25.7356 | -35.5341 | -35.8935
25.7417 | -35.5344 | -35.899
25.7478 | -35.5347 | -35.9042
25.7539 | -35.535 | -35.9097
25.76 | -35.5354 | -35.9155
25.7661 | -35.5357 | -35.921
25.7722 | -35.5361 | -35.9266
25.7783 | -35.5364 | -35.9321
25.7844 | -35.5368 | -35.9375
25.7905 | -35.5371 | -35.9429
25.7966 | -35.5374 | -35.9482
25.8027 | -35.5377 | -35.9538
25.8088 | -35.538 | -35.9597
25.8149 | -35.5383 | -35.9654
25.821 | -35.5386 | -35.9708
25.8271 | -35.5389 | -35.9765
25.8332 | -35.5393 | -35.9819
25.8393 | -35.5397 | -35.9874
25.8454 | -35.54 | -35.9924
25.8515 | -35.5404 | -35.9978
25.8576 | -35.5408 | -36.003
25.8638 | -35.5411 | -36.0085
25.8699 | -35.5415 | -36.0144
25.876 | -35.5418 | -36.0197
25.8821 | -35.5421 | -36.0254
25.8882 | -35.5424 | -36.0307
25.8943 | -35.5428 | -36.0363
25.9004 | -35.543 | -36.0418
25.9065 | -35.5434 | -36.047
25.9126 | -35.5438 | -36.0526
25.9187 | -35.5441 | -36.058
25.9248 | -35.5445 | -36.0635
25.9309 | -35.5448 | -36.0691
25.937 | -35.5452 | -36.0748
25.9431 | -35.5456 | -36.0804
25.9492 | -35.5459 | -36.0859
25.9553 | -35.5462 | -36.0917
25.9614 | -35.5467 | -36.0967
25.9675 | -35.547 | -36.1022
25.9736 | -35.5473 | -36.1079
25.9797 | -35.5476 | -36.1137
25.9858 | -35.548 | -36.119
25.9919 | -35.5483 | -36.1247
25.998 | -35.5487 | -36.13
26.0042 | -35.549 | -36.1353
26.0103 | -35.5493 | -36.1407
26.0164 | -35.5496 | -36.1463
26.0225 | -35.5499 | -36.1515
26.0286 | -35.5503 | -36.1571
26.0347 | -35.5506 | -36.1623
26.0408 | -35.551 | -36.168
26.0469 | -35.5513 | -36.1732
26.053 | -35.5515 | -36.1789
26.0591 | -35.5519 | -36.1844
26.0652 | -35.5523 | -36.19
26.0713 | -35.5527 | -36.1954
26.0774 | -35.553 | -36.2009
26.0835 | -35.5534 | -36.2063
26.0896 | -35.5537 | -36.2118
26.0957 | -35.5541 | -36.2173
26.1018 | -35.5544 | -36.2226
26.1079 | -35.5548 | -36.2279
26.114 | -35.555 | -36.2335
26.1201 | -35.5553 | -36.2391
26.1262 | -35.5556 | -36.2448
26.1323 | -35.5559 | -36.25
26.1384 | -35.5563 | -36.2551
26.1445 | -35.5567 | -36.2604
26.1507 | -35.557 | -36.2657
26.1568 | -35.5573 | -36.2714
26.1629 | -35.5576 | -36.2767
26.169 | -35.5579 | -36.2822
26.1751 | -35.5582 | -36.2874
26.1812 | -35.5585 | -36.2931
26.1873 | -35.5588 | -36.2988
26.1934 | -35.5592 | -36.3041
26.1995 | -35.5595 | -36.3094
26.2056 | -35.5598 | -36.3145
26.2117 | -35.5602 | -36.3198
26.2178 | -35.5605 | -36.3252
26.2239 | -35.5609 | -36.3304
26.23 | -35.5612 | -36.3359
26.2361 | -35.5615 | -36.3411
26.2422 | -35.5618 | -36.3464
26.2483 | -35.5621 | -36.3517
26.2544 | -35.5625 | -36.3568
26.2605 | -35.5628 | -36.362
26.2666 | -35.5631 | -36.3673
26.2727 | -35.5634 | -36.3729
26.2788 | -35.5638 | -36.3779
26.2849 | -35.5641 | -36.3834
26.2911 | -35.5644 | -36.3885
26.2972 | -35.5647 | -36.3942
26.3033 | -35.5651 | -36.3994
26.3094 | -35.5654 | -36.405
26.3155 | -35.5658 | -36.4099
26.3216 | -35.5661 | -36.4153
26.3277 | -35.5664 | -36.4204
26.3338 | -35.5668 | -36.4256
26.3399 | -35.5671 | -36.431
26.346 | -35.5674 | -36.436
26.3521 | -35.5677 | -36.4417
26.3582 | -35.568 | -36.4471
26.3643 | -35.5684 | -36.4525
26.3704 | -35.5687 | -36.458
26.3765 | -35.5691 | -36.4633
26.3826 | -35.5694 | -36.4684
26.3887 | -35.5698 | -36.4736
26.3948 | -35.5701 | -36.4792
26.4009 | -35.5704 | -36.4844
26.407 | -35.5708 | -36.4896
26.4131 | -35.5711 | -36.4949
26.4192 | -35.5714 | -36.5001
26.4253 | -35.5717 | -36.5053
26.4314 | -35.572 | -36.5104
26.4376 | -35.5724 | -36.516
26.4437 | -35.5726 | -36.521
26.4498 | -35.5729 | -36.5262
26.4559 | -35.5732 | -36.531
26.462 | -35.5735 | -36.5366
26.4681 | -35.5739 | -36.5419
26.4742 | -35.5742 | -36.5475
26.4803 | -35.5746 | -36.5527
26.4864 | -35.5749 | -36.5581
26.4925 | -35.5753 | -36.5634
26.4986 | -35.5756 | -36.5689
26.5047 | -35.5759 | -36.5741
26.5108 | -35.5763 | -36.5793
26.5169 | -35.5766 | -36.5849
26.523 | -35.5769 | -36.5902
26.5291 | -35.5772 | -36.5952
26.5352 | -35.5774 | -36.6008
26.5413 | -35.5778 | -36.6062
26.5474 | -35.5781 | -36.6115
26.5535 | -35.5784 | -36.6168
26.5596 | -35.5787 | -36.6219
26.5657 | -35.579 | -36.627
26.5718 | -35.5793 | -36.6327
26.578 | -35.5796 | -36.6381
26.5841 | -35.58 | -36.6436
26.5902 | -35.5803 | -36.6488
26.5963 | -35.5806 | -36.6538
26.6024 | -35.5808 | -36.659
26.6085 | -35.5811 | -36.6644
26.6146 | -35.5813 | -36.6696
26.6207 | -35.5817 | -36.6751
26.6268 | -35.582 | -36.6804
26.6329 | -35.5823 | -36.6859
26.639 | -35.5826 | -36.6912
26.6451 | -35.5829 | -36.6966
26.6512 | -35.5832 | -36.7018
26.6573 | -35.5835 | -36.7074
26.6634 | -35.5838 | -36.7122
26.6695 | -35.5842 | -36.7175
26.6756 | -35.5845 | -36.7231
26.6817 | -35.5848 | -36.7284
26.6878 | -35.5851 | -36.7338
26.6939 | -35.5854 | -36.739
26.7 | -35.5857 | -36.744
26.7061 | -35.586 | -36.7492
26.7122 | -35.5864 | -36.7545
26.7183 | -35.5867 | -36.7598
26.7245 | -35.587 | -36.7653
26.7306 | -35.5873 | -36.7704
26.7367 | -35.5877 | -36.7755
26.7428 | -35.588 | -36.7806
26.7489 | -35.5883 | -36.7862
26.755 | -35.5886 | -36.7915
26.7611 | -35.589 | -36.7966
26.7672 | -35.5893 | -36.802
26.7733 | -35.5896 | -36.8071
26.7794 | -35.5899 | -36.8124
26.7855 | -35.5902 | -36.8177
26.7916 | -35.5905 | -36.8227
26.7977 | -35.5908 | -36.8279
26.8038 | -35.5912 | -36.8329
26.8099 | -35.5915 | -36.838
26.816 | -35.5918 | -36.843
26.8221 | -35.5921 | -36.8479
26.8282 | -35.5924 | -36.8532
26.8343 | -35.5928 | -36.8585
26.8404 | -35.5931 | -36.8635
26.8465 | -35.5934 | -36.8687
26.8526 | -35.5937 | -36.874
26.8587 | -35.594 | -36.8793
26.8649 | -35.5943 | -36.8845
26.871 | -35.5946 | -36.8899
26.8771 | -35.5949 | -36.8947
26.8832 | -35.5952 | -36.9001
26.8893 | -35.5955 | -36.9054
26.8954 | -35.5959 | -36.9107
26.9015 | -35.5962 | -36.9161
26.9076 | -35.5965 | -36.9214
26.9137 | -35.5969 | -36.9263
26.9198 | -35.5972 | -36.9315
26.9259 | -35.5975 | -36.9364
26.932 | -35.5978 | -36.9413
26.9381 | -35.5981 | -36.9469
26.9442 | -35.5984 | -36.9515
26.9503 | -35.5987 | -36.9571
26.9564 | -35.599 | -36.9625
26.9625 | -35.5993 | -36.9677
26.9686 | -35.5997 | -36.9733
26.9747 | -35.6001 | -36.9786
26.9808 | -35.6004 | -36.9836
26.9869 | -35.6007 | -36.9889
26.993 | -35.601 | -36.9941
26.9991 | -35.6013 | -36.999
27.0052 | -35.6016 | -37.0045
27.0114 | -35.6019 | -37.01
27.0175 | -35.6022 | -37.015
27.0236 | -35.6025 | -37.02
27.0297 | -35.6028 | -37.0254
27.0358 | -35.6031 | -37.0305
27.0419 | -35.6035 | -37.036
27.048 | -35.6038 | -37.0411
27.0541 | -35.6041 | -37.0464
27.0602 | -35.6044 | -37.0515
27.0663 | -35.6047 | -37.0568
27.0724 | -35.605 | -37.0618
27.0785 | -35.6053 | -37.0671
27.0846 | -35.6056 | -37.072
27.0907 | -35.606 | -37.0771
27.0968 | -35.6063 | -37.0823
27.1029 | -35.6066 | -37.0873
27.109 | -35.607 | -37.0926
27.1151 | -35.6073 | -37.0976
27.1212 | -35.6075 | -37.1025
27.1273 | -35.6079 | -37.1079
27.1334 | -35.6081 | -37.1133
27.1395 | -35.6085 | -37.1185
27.1456 | -35.6088 | -37.1239
27.1518 | -35.6091 | -37.129
27.1579 | -35.6094 | -37.1339
27.164 | -35.6097 | -37.139
27.1701 | -35.61 | -37.1442
27.1762 | -35.6103 | -37.1494
27.1823 | -35.6106 | -37.1544
27.1884 | -35.6109 | -37.1597
27.1945 | -35.6112 | -37.1648
27.2006 | -35.6116 | -37.1698
27.2067 | -35.6119 | -37.1748
27.2128 | -35.6122 | -37.18
27.2189 | -35.6125 | -37.1852
27.225 | -35.6128 | -37.1905
27.2311 | -35.6131 | -37.1952
27.2372 | -35.6134 | -37.2005
27.2433 | -35.6137 | -37.2056
27.2494 | -35.614 | -37.2107
27.2555 | -35.6144 | -37.2161
27.2616 | -35.6146 | -37.221
27.2677 | -35.615 | -37.2261
27.2738 | -35.6153 | -37.2311
27.2799 | -35.6156 | -37.2364
27.286 | -35.6159 | -37.2414
27.2921 | -35.6162 | -37.2465
27.2983 | -35.6165 | -37.2516
27.3044 | -35.6168 | -37.2567
27.3105 | -35.6171 | -37.2621
27.3166 | -35.6175 | -37.2671
27.3227 | -35.6178 | -37.2721
27.3288 | -35.6181 | -37.2771
27.3349 | -35.6185 | -37.2822
27.341 | -35.6188 | -37.2874
27.3471 | -35.6191 | -37.2924
27.3532 | -35.6194 | -37.2975
27.3593 | -35.6197 | -37.3027
27.3654 | -35.6201 | -37.3081
27.3715 | -35.6204 | -37.313
27.3776 | -35.6207 | -37.3183
27.3837 | -35.621 | -37.3233
27.3898 | -35.6213 | -37.3284
27.3959 | -35.6216 | -37.3334
27.402 | -35.622 | -37.3384
27.4081 | -35.6223 | -37.3435
27.4142 | -35.6226 | -37.3485
27.4203 | -35.623 | -37.3535
27.4264 | -35.6232 | -37.3586
27.4325 | -35.6235 | -37.3637
27.4387 | -35.6239 | -37.3687
27.4448 | -35.6242 | -37.374
27.4509 | -35.6245 | -37.3792
27.457 | -35.6248 | -37.3843
27.4631 | -35.6251 | -37.3894
27.4692 | -35.6254 | -37.3944
27.4753 | -35.6257 | -37.3991
27.4814 | -35.626 | -37.4041
27.4875 | -35.6263 | -37.4095
27.4936 | -35.6266 | -37.4147
27.4997 | -35.6269 | -37.4197
27.5058 | -35.6272 | -37.425
27.5119 | -35.6275 | -37.43
27.518 | -35.6279 | -37.4351
27.5241 | -35.6282 | -37.4403
27.5302 | -35.6285 | -37.4456
27.5363 | -35.6288 | -37.4507
27.5424 | -35.6291 | -37.4556
27.5485 | -35.6295 | -37.4608
27.5546 | -35.6298 | -37.4657
27.5607 | -35.6301 | -37.4709
27.5668 | -35.6304 | -37.4758
27.5729 | -35.6307 | -37.4809
27.5791 | -35.631 | -37.4858
27.5852 | -35.6313 | -37.4909
27.5913 | -35.6316 | -37.4963
27.5974 | -35.6319 | -37.5015
27.6035 | -35.6322 | -37.5066
27.6096 | -35.6325 | -37.5114
27.6157 | -35.6328 | -37.5164
27.6218 | -35.6331 | -37.5215
27.6279 | -35.6334 | -37.5268
27.634 | -35.6338 | -37.532
27.6401 | -35.634 | -37.5371
27.6462 | -35.6343 | -37.5421
27.6523 | -35.6347 | -37.5471
27.6584 | -35.6349 | -37.5522
27.6645 | -35.6353 | -37.5575
27.6706 | -35.6356 | -37.5624
27.6767 | -35.6359 | -37.5676
27.6828 | -35.6362 | -37.5724
27.6889 | -35.6365 | -37.5775
27.695 | -35.6368 | -37.5828
27.7011 | -35.6372 | -37.5881
27.7072 | -35.6375 | -37.5933
27.7133 | -35.6379 | -37.5985
27.7194 | -35.6382 | -37.6034
27.7256 | -35.6384 | -37.6084
27.7317 | -35.6387 | -37.6134
27.7378 | -35.639 | -37.6186
27.7439 | -35.6393 | -37.6238
27.75 | -35.6396 | -37.6291
27.7561 | -35.6399 | -37.6339
27.7622 | -35.6402 | -37.639
27.7683 | -35.6405 | -37.6443
27.7744 | -35.6408 | -37.6496
27.7805 | -35.6411 | -37.6548
27.7866 | -35.6414 | -37.6599
27.7927 | -35.6417 | -37.665
27.7988 | -35.642 | -37.67
27.8049 | -35.6423 | -37.6749
27.811 | -35.6426 | -37.68
27.8171 | -35.6429 | -37.6851
27.8232 | -35.6432 | -37.6899
27.8293 | -35.6435 | -37.6949
27.8354 | -35.6439 | -37.7002
27.8415 | -35.6442 | -37.7054
27.8476 | -35.6445 | -37.7105
27.8537 | -35.6448 | -37.7157
27.8598 | -35.6451 | -37.7207
27.866 | -35.6454 | -37.7256
27.8721 | -35.6457 | -37.7306
27.8782 | -35.646 | -37.7355
27.8843 | -35.6463 | -37.7408
27.8904 | -35.6466 | -37.7458
27.8965 | -35.6469 | -37.7512
27.9026 | -35.6472 | -37.7563
27.9087 | -35.6475 | -37.7614
27.9148 | -35.6478 | -37.7665
27.9209 | -35.6481 | -37.7719
27.927 | -35.6484 | -37.7768
27.9331 | -35.6487 | -37.7822
27.9392 | -35.649 | -37.7873
27.9453 | -35.6493 | -37.7924
27.9514 | -35.6496 | -37.7973
27.9575 | -35.6499 | -37.8023
27.9636 | -35.6502 | -37.8074
27.9697 | -35.6505 | -37.8122
27.9758 | -35.6508 | -37.8172
27.9819 | -35.6511 | -37.8224
27.988 | -35.6514 | -37.8276
27.9941 | -35.6517 | -37.8326
28.0002 | -35.652 | -37.8375
28.0063 | -35.6523 | -37.8428
28.0125 | -35.6526 | -37.8478
28.0186 | -35.6529 | -37.853
28.0247 | -35.6532 | -37.858
28.0308 | -35.6535 | -37.8631
28.0369 | -35.6538 | -37.8679
28.043 | -35.6541 | -37.8731
28.0491 | -35.6544 | -37.8781
28.0552 | -35.6547 | -37.8833
28.0613 | -35.6549 | -37.8884
28.0674 | -35.6553 | -37.8934
28.0735 | -35.6555 | -37.8983

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.36 | -42.0953
21.4908 | -35.4271 | -34.846
16.4948 | -34.8117 | -27.4918
12.6602 | -34.4606 | -21.2453
9.7171 | -34.2541 | -16.3576
7.4582 | -34.1078 | -12.414
5.7244 | -34.0059 | -9.2163
4.3936 | -33.9397 | -6.6321
3.3722 | -33.9292 | -4.5973
2.5883 | -33.9234 | -3.0549
1.9866 | -33.9292 | -1.8203
1.5248 | -33.936 | -0.83057
1.1703 | -33.9536 | -0.18969
1 | -33.9606 | 0.12699
0.89822 | -33.9597 | 0.40137
0.68941 | -33.9566 | 0.9766
0.52919 | -33.9672 | 1.4244
0.40612 | -33.9599 | 2.0317
0.31171 | -33.9429 | 2.4833
0.23928 | -33.9523 | 3.2142
0.18358 | -33.9628 | 3.6376
0.14091 | -33.9668 | 4.528
0.10819 | -33.9722 | 5.6978
0.083017 | -34.0177 | 7.1793
0.063705 | -34.0622 | 9.058
0.048923 | -34.1797 | 11.5618
0.037479 | -34.2854 | 14.7506
0.028801 | -34.4637 | 18.7782
0.022078 | -34.7392 | 24.1065
0.017023 | -35.2253 | 30.5112
0.013018 | -36.0223 | 38.1976
0.010014 | -37.1438 | 46.0945

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003727 | -4.5088
0.007454 | -25.0599
0.011181 | -33.8877
0.014908 | -39.7622
0.018635 | -42.7663
0.022362 | -43.6569
0.026089 | -46.5946
0.029816 | -50.1691
0.033543 | -49.4983
0.03727 | -50.0798
0.040997 | -51.7642
0.044724 | -53.6627
0.048451 | -54.7554
0.052178 | -55.253
0.055905 | -57.7745
0.059632 | -57.9154
0.063359 | -57.1683
0.067086 | -57.9825
0.070813 | -59.195
0.07454 | -59.8017
0.078267 | -60.3586
0.081994 | -60.5745
0.085721 | -61.2147
0.089448 | -61.6362
0.093175 | -61.3498
0.096902 | -62.6649
0.10063 | -62.929
0.10436 | -63.1853
0.10808 | -63.1846
0.11181 | -63.9942
0.11554 | -65.0582
0.11926 | -65.0586
0.12299 | -66.7094
0.12672 | -66.8786
0.13045 | -66.6379
0.13417 | -67.3708
0.1379 | -67.0815
0.14163 | -67.5725
0.14535 | -67.9397
0.14908 | -67.5198
0.15281 | -67.7139
0.15653 | -68.3236
0.16026 | -68.2788
0.16399 | -69.1729
0.16772 | -69.6466
0.17144 | -69.2503
0.17517 | -68.5143
0.1789 | -69.4181
0.18262 | -71.5659
0.18635 | -71.1893
0.19008 | -69.932
0.1938 | -69.8567
0.19753 | -70.7251
0.20126 | -71.187
0.20499 | -71.7321
0.20871 | -71.5494
0.21244 | -71.5253
0.21617 | -71.9624
0.21989 | -71.6648
0.22362 | -71.9256
0.22735 | -71.8691
0.23107 | -72.6167
0.2348 | -72.979
0.23853 | -74.204
0.24225 | -74.2024
0.24598 | -74.1133
0.24971 | -73.2958
0.25344 | -73.3637
0.25716 | -73.4727
0.26089 | -73.4833
0.26462 | -73.72
0.26834 | -73.2067
0.27207 | -73.8122
0.2758 | -73.8102
0.27953 | -73.8741
0.28325 | -73.6636
0.28698 | -74.0587
0.29071 | -74.3046
0.29443 | -73.7239
0.29816 | -74.863
0.30189 | -75.0239
0.30561 | -74.2522
0.30934 | -75.0774
0.31307 | -75.6963
0.31679 | -75.6423
0.32052 | -77.0632
0.32425 | -76.5725
0.32798 | -75.2981
0.3317 | -75.2291
0.33543 | -74.3989
0.33916 | -74.577
0.34288 | -75.2534
0.34661 | -75.3755
0.35034 | -75.4556
0.35407 | -75.3019
0.35779 | -76.8522
0.36152 | -76.7998
0.36525 | -76.6347
0.36897 | -77.4312
0.3727 | -76.8111
0.37643 | -76.1397
0.38015 | -76.5139
0.38388 | -77.3478
0.38761 | -77.396
0.39133 | -77.5991
0.39506 | -77.4355
0.39879 | -77.8415
0.40252 | -77.2377
0.40624 | -76.883
0.40997 | -76.9561
0.4137 | -78.0854
0.41742 | -78.1032
0.42115 | -78.2149
0.42488 | -77.7486
0.42861 | -77.0675
0.43233 | -77.7122
0.43606 | -78.2922
0.43979 | -78.3824
0.44351 | -78.5994
0.44724 | -78.3672
0.45097 | -78.7687
0.45469 | -78.4507
0.45842 | -78.7731
0.46215 | -79.6045
0.46587 | -79.7121
0.4696 | -78.5369
0.47333 | -78.7799
0.47706 | -79.4355
0.48078 | -79.5534
0.48451 | -78.4276
0.48824 | -79.4261
0.49196 | -80.8941
0.49569 | -80.3485
0.49942 | -79.71
0.50314 | -80.0567
0.50687 | -79.9757
0.5106 | -80.8773
0.51433 | -80.2837
0.51805 | -79.8985
0.52178 | -79.3836
0.52551 | -79.2757
0.52923 | -79.1868
0.53296 | -79.0153
0.53669 | -79.7117
0.54041 | -80.382
0.54414 | -80.4037
0.54787 | -80.153
0.5516 | -80.0955
0.55532 | -80.1428
0.55905 | -79.9913
0.56278 | -79.7561
0.5665 | -80.0497
0.57023 | -80.7752
0.57396 | -81.27
0.57769 | -81.0049
0.58141 | -81.0012
0.58514 | -81.3391
0.58887 | -81.3634
0.59259 | -81.4741
0.59632 | -81.6165
0.60005 | -81.1407
0.60377 | -80.6716
0.6075 | -80.9694
0.61123 | -81.2278
0.61496 | -80.6978
0.61868 | -80.6691
0.62241 | -81.0333
0.62614 | -81.2539
0.62986 | -81.3459
0.63359 | -81.3933
0.63732 | -81.5455
0.64104 | -81.4135
0.64477 | -81.4475
0.6485 | -82.198
0.65223 | -82.9621
0.65595 | -82.6304
0.65968 | -82.151
0.66341 | -82.211
0.66713 | -82.246
0.67086 | -81.7204
0.67459 | -81.6384
0.67831 | -81.8346
0.68204 | -82.0486
0.68577 | -82.5112
0.68949 | -82.9626
0.69322 | -83.5333
0.69695 | -84.0062
0.70068 | -83.7467
0.7044 | -83.2822
0.70813 | -83.3003
0.71186 | -83.5627
0.71558 | -83.6519
0.71931 | -83.4515
0.72304 | -83.0748
0.72676 | -82.7789
0.73049 | -82.698
0.73422 | -82.526
0.73794 | -82.28
0.74167 | -82.5423
0.7454 | -83.4328
0.74913 | -83.9755
0.75285 | -83.6286
0.75658 | -83.0185
0.76031 | -82.7522
0.76403 | -82.5179
0.76776 | -82.724
0.77149 | -83.1051
0.77521 | -82.9168
0.77894 | -82.8587
0.78267 | -82.9825
0.7864 | -83.0965
0.79012 | -83.244
0.79385 | -83.5163
0.79758 | -83.9325
0.8013 | -84.3033
0.80503 | -84.6095
0.80876 | -84.8706
0.81249 | -84.319
0.81621 | -83.3411
0.81994 | -83.2802
0.82367 | -83.6684
0.82739 | -83.7505
0.83112 | -83.8394
0.83485 | -84.2558
0.83857 | -84.4154
0.8423 | -84.0575
0.84603 | -83.7976
0.84976 | -84.1377
0.85348 | -84.6428
0.85721 | -84.6576
0.86094 | -84.2559
0.86466 | -84.3238
0.86839 | -84.5364
0.87212 | -84.46
0.87584 | -84.1865
0.87957 | -83.9537
0.8833 | -84.2517
0.88702 | -84.8762
0.89075 | -85.2989
0.89448 | -85.4442
0.89821 | -85.4592
0.90193 | -85.4918
0.90566 | -85.4015
0.90939 | -85.0184
0.91311 | -84.6541
0.91684 | -84.7654
0.92057 | -84.9953
0.92429 | -85.3919
0.92802 | -85.9828
0.93175 | -86.1582
0.93548 | -85.9214
0.9392 | -85.4147
0.94293 | -84.7859
0.94666 | -84.3909
0.95038 | -84.5173
0.95411 | -84.7633
0.95784 | -84.8177
0.96157 | -85.1765
0.96529 | -85.4857
0.96902 | -85.9186
0.97275 | -86.2423
0.97647 | -85.8732
0.9802 | -85.4076
0.98393 | -85.3505
0.98765 | -86.0499
0.99138 | -86.0265
0.99511 | -85.1849
0.99884 | -85.0306
1.0026 | -85.3903
1.0063 | -85.8995
1.01 | -86.2098
1.0137 | -86.0372
1.0175 | -85.5388
1.0212 | -85.4636
1.0249 | -85.7968
1.0287 | -86.1531
1.0324 | -86.8219
1.0361 | -87.1013
1.0398 | -86.2367
1.0436 | -85.6422
1.0473 | -85.9657
1.051 | -86.2606
1.0547 | -86.4122
1.0585 | -86.7771
1.0622 | -87.2864
1.0659 | -87.3497
1.0696 | -87.163
1.0734 | -87.1684
1.0771 | -87.0265
1.0808 | -86.7759
1.0846 | -86.503
1.0883 | -86.3041
1.092 | -86.2832
1.0957 | -86.5046
1.0995 | -86.8918
1.1032 | -87.1963
1.1069 | -87.2804
1.1106 | -87.1769
1.1144 | -86.9215
1.1181 | -86.6979
1.1218 | -86.6455
1.1256 | -86.6775
1.1293 | -86.7158
1.133 | -86.7265
1.1367 | -86.7585
1.1405 | -86.8787
1.1442 | -87.0388
1.1479 | -87.0907
1.1516 | -86.9366
1.1554 | -86.7219
1.1591 | -86.6456
1.1628 | -86.7107
1.1665 | -86.8398
1.1703 | -87.0042
1.174 | -87.2701
1.1777 | -87.577
1.1815 | -87.6563
1.1852 | -87.5328
1.1889 | -87.5158
1.1926 | -87.6959
1.1964 | -87.9174
1.2001 | -88.0402
1.2038 | -88.0101
1.2075 | -87.8504
1.2113 | -87.6552
1.215 | -87.5251
1.2187 | -87.4681
1.2225 | -87.4995
1.2262 | -87.6509
1.2299 | -87.8764
1.2336 | -88.0743
1.2374 | -88.1992
1.2411 | -88.1801
1.2448 | -88.0165
1.2485 | -87.8013
1.2523 | -87.5971
1.256 | -87.488
1.2597 | -87.5686
1.2635 | -87.8225
1.2672 | -88.0926
1.2709 | -88.27
1.2746 | -88.2861
1.2784 | -88.0815
1.2821 | -87.9212
1.2858 | -87.9478
1.2895 | -87.9463
1.2933 | -87.9226
1.297 | -88.0461
1.3007 | -88.2146
1.3044 | -88.2455
1.3082 | -88.2715
1.3119 | -88.4879
1.3156 | -88.7339
1.3194 | -88.8794
1.3231 | -89.1528
1.3268 | -89.5458
1.3305 | -89.5669
1.3343 | -89.1893
1.338 | -88.9689
1.3417 | -89.0773
1.3454 | -89.2043
1.3492 | -89.0611
1.3529 | -88.7354
1.3566 | -88.4919
1.3604 | -88.437
1.3641 | -88.527
1.3678 | -88.6931
1.3715 | -88.798
1.3753 | -88.6262
1.379 | -88.2178
1.3827 | -87.9847
1.3864 | -88.2133
1.3902 | -88.7169
1.3939 | -89.0119
1.3976 | -88.8917
1.4014 | -88.6533
1.4051 | -88.6845
1.4088 | -89.0164
1.4125 | -89.3297
1.4163 | -89.511
1.42 | -89.7273
1.4237 | -89.8757
1.4274 | -89.7968
1.4312 | -89.6904
1.4349 | -89.7563
1.4386 | -89.8778
1.4423 | -89.8603
1.4461 | -89.8177
1.4498 | -89.9737
1.4535 | -90.1614
1.4573 | -89.9752
1.461 | -89.3737
1.4647 | -88.7724
1.4684 | -88.5736
1.4722 | -88.7976
1.4759 | -89.2053
1.4796 | -89.5643
1.4833 | -89.6612
1.4871 | -89.4308
1.4908 | -89.1141
1.4945 | -88.9646
1.4983 | -89.0266
1.502 | -89.0847
1.5057 | -88.9882
1.5094 | -88.9371
1.5132 | -89.1711
1.5169 | -89.5507
1.5206 | -89.7188
1.5243 | -89.7453
1.5281 | -89.9976
1.5318 | -90.4325
1.5355 | -90.6172
1.5393 | -90.3655
1.543 | -89.9725
1.5467 | -89.7522
1.5504 | -89.7027
1.5542 | -89.7238
1.5579 | -89.9079
1.5616 | -90.1957
1.5653 | -90.2811
1.5691 | -90.1639
1.5728 | -90.1242
1.5765 | -90.2505
1.5802 | -90.2135
1.584 | -89.7836
1.5877 | -89.2904
1.5914 | -89.0878
1.5952 | -89.216
1.5989 | -89.4544
1.6026 | -89.6045
1.6063 | -89.6722
1.6101 | -89.642
1.6138 | -89.6231
1.6175 | -89.6287
1.6212 | -89.7043
1.625 | -89.8228
1.6287 | -89.8971
1.6324 | -89.8897
1.6362 | -89.8161
1.6399 | -89.721
1.6436 | -89.6435
1.6473 | -89.6536
1.6511 | -89.795
1.6548 | -90.0146
1.6585 | -90.1839
1.6622 | -90.2161
1.666 | -90.1597
1.6697 | -90.1044
1.6734 | -90.0471
1.6771 | -89.9569
1.6809 | -89.827
1.6846 | -89.7203
1.6883 | -89.7064
1.6921 | -89.8148
1.6958 | -90.0134
1.6995 | -90.2008
1.7032 | -90.2734
1.707 | -90.2006
1.7107 | -90.1025
1.7144 | -90.1283
1.7181 | -90.2871
1.7219 | -90.4977
1.7256 | -90.725
1.7293 | -90.9726
1.7331 | -91.1995
1.7368 | -91.322
1.7405 | -91.3076
1.7442 | -91.2099
1.748 | -91.0737
1.7517 | -90.9548
1.7554 | -90.9037
1.7591 | -90.9525
1.7629 | -91.0376
1.7666 | -91.0571
1.7703 | -91.044
1.7741 | -91.094
1.7778 | -91.1725
1.7815 | -91.1364
1.7852 | -90.9509
1.789 | -90.8187
1.7927 | -90.9092
1.7964 | -91.2044
1.8001 | -91.5295
1.8039 | -91.6804
1.8076 | -91.6026
1.8113 | -91.4117
1.815 | -91.2717
1.8188 | -91.2342
1.8225 | -91.2215
1.8262 | -91.1841
1.83 | -91.1587
1.8337 | -91.2001
1.8374 | -91.2846
1.8411 | -91.3148
1.8449 | -91.2348
1.8486 | -91.117
1.8523 | -91.0317
1.856 | -90.9899
1.8598 | -90.9792
1.8635 | -90.9944
1.8672 | -91.033
1.871 | -91.084
1.8747 | -91.1051
1.8784 | -91.0696
1.8821 | -91.0155
1.8859 | -91.0675
1.8896 | -91.2778
1.8933 | -91.554
1.897 | -91.7664
1.9008 | -91.8721
1.9045 | -91.8957
1.9082 | -91.8535
1.912 | -91.7426
1.9157 | -91.6178
1.9194 | -91.521
1.9231 | -91.4192
1.9269 | -91.285
1.9306 | -91.1511
1.9343 | -91.0319
1.938 | -90.9496
1.9418 | -90.9905
1.9455 | -91.243
1.9492 | -91.6213
1.9529 | -91.9476
1.9567 | -92.1378
1.9604 | -92.2255
1.9641 | -92.2627
1.9679 | -92.2784
1.9716 | -92.2705
1.9753 | -92.2515
1.979 | -92.2649
1.9828 | -92.3145
1.9865 | -92.3109
1.9902 | -92.2006
1.9939 | -92.0622
1.9977 | -91.9875
2.0014 | -92.0097
2.0051 | -92.1178
2.0089 | -92.2979
2.0126 | -92.4984
2.0163 | -92.6013
2.02 | -92.532
2.0238 | -92.3824
2.0275 | -92.2575
2.0312 | -92.1847
2.0349 | -92.0808
2.0387 | -91.9184
2.0424 | -91.7906
2.0461 | -91.7885
2.0498 | -91.8946
2.0536 | -92.0016
2.0573 | -92.0635
2.061 | -92.1347
2.0648 | -92.2329
2.0685 | -92.3414
2.0722 | -92.4421
2.0759 | -92.5249
2.0797 | -92.5785
2.0834 | -92.6042
2.0871 | -92.6244
2.0908 | -92.6375
2.0946 | -92.634
2.0983 | -92.6087
2.102 | -92.5922
2.1058 | -92.6343
2.1095 | -92.7606
2.1132 | -92.9132
2.1169 | -93.034
2.1207 | -93.0826
2.1244 | -93.0784
2.1281 | -93.0738
2.1318 | -93.1419
2.1356 | -93.2774
2.1393 | -93.3987
2.143 | -93.4258
2.1468 | -93.3631
2.1505 | -93.2205
2.1542 | -93.0875
2.1579 | -92.934
2.1617 | -92.7834
2.1654 | -92.6622
2.1691 | -92.6031
2.1728 | -92.6171
2.1766 | -92.6755
2.1803 | -92.7246
2.184 | -92.7413
2.1877 | -92.7505
2.1915 | -92.7864
2.1952 | -92.8785
2.1989 | -93.0148
2.2027 | -93.1368
2.2064 | -93.2089
2.2101 | -93.2091
2.2138 | -93.1631
2.2176 | -93.123
2.2213 | -93.129
2.225 | -93.1879
2.2287 | -93.2911
2.2325 | -93.391
2.2362 | -93.4231
2.2399 | -93.372
2.2437 | -93.279
2.2474 | -93.196
2.2511 | -93.1566
2.2548 | -93.1504
2.2586 | -93.1357
2.2623 | -93.0703
2.266 | -92.9433
2.2697 | -92.783
2.2735 | -92.6371
2.2772 | -92.5571
2.2809 | -92.5633
2.2846 | -92.6377
2.2884 | -92.7348
2.2921 | -92.7964
2.2958 | -92.8295
2.2996 | -92.87
2.3033 | -92.94
2.307 | -93.0352
2.3107 | -93.12
2.3145 | -93.1554
2.3182 | -93.1309
2.3219 | -93.107
2.3256 | -93.1681
2.3294 | -93.3782
2.3331 | -93.7127
2.3368 | -94.0383
2.3406 | -94.2154
2.3443 | -94.2077
2.348 | -94.0612
2.3517 | -93.8698
2.3555 | -93.691
2.3592 | -93.5337
2.3629 | -93.4104
2.3666 | -93.3541
2.3704 | -93.3857
2.3741 | -93.4861
2.3778 | -93.5826
2.3816 | -93.5957
2.3853 | -93.4828
2.389 | -93.315
2.3927 | -93.1888
2.3965 | -93.1487
2.4002 | -93.1867
2.4039 | -93.2546
2.4076 | -93.2911
2.4114 | -93.2627
2.4151 | -93.1864
2.4188 | -93.1182
2.4225 | -93.0867
2.4263 | -93.1078
2.43 | -93.1805
2.4337 | -93.274
2.4375 | -93.3549
2.4412 | -93.3916
2.4449 | -93.3761
2.4486 | -93.3416
2.4524 | -93.3061
2.4561 | -93.2625
2.4598 | -93.1944
2.4635 | -93.1137
2.4673 | -93.0323
2.471 | -92.9943
2.4747 | -93.0731
2.4785 | -93.28
2.4822 | -93.5684
2.4859 | -93.8342
2.4896 | -93.9834
2.4934 | -94.0222
2.4971 | -94.0323
2.5008 | -94.0737
2.5045 | -94.1484
2.5083 | -94.2117
2.512 | -94.2146
2.5157 | -94.15
2.5194 | -94.0677
2.5232 | -94.0275
2.5269 | -94.0678
2.5306 | -94.1452
2.5344 | -94.2095
2.5381 | -94.2497
2.5418 | -94.2774
2.5455 | -94.2954
2.5493 | -94.3147
2.553 | -94.352
2.5567 | -94.4451
2.5604 | -94.601
2.5642 | -94.7661
2.5679 | -94.8825
2.5716 | -94.9111
2.5754 | -94.8334
2.5791 | -94.6536
2.5828 | -94.3982
2.5865 | -94.1124
2.5903 | -93.8379
2.594 | -93.6396
2.5977 | -93.5452
2.6014 | -93.5503
2.6052 | -93.6313
2.6089 | -93.7526
2.6126 | -93.8767
2.6164 | -93.979
2.6201 | -94.0313
2.6238 | -94.0112
2.6275 | -93.9425
2.6313 | -93.8638
2.635 | -93.8179
2.6387 | -93.8272
2.6424 | -93.8613
2.6462 | -93.88
2.6499 | -93.8687
2.6536 | -93.8445
2.6573 | -93.8089
2.6611 | -93.7652
2.6648 | -93.7089
2.6685 | -93.6389
2.6723 | -93.5966
2.676 | -93.605
2.6797 | -93.652
2.6834 | -93.7251
2.6872 | -93.7549
2.6909 | -93.7832
2.6946 | -93.8093
2.6983 | -93.8564
2.7021 | -93.9531
2.7058 | -94.1068
2.7095 | -94.3052
2.7133 | -94.5012
2.717 | -94.654
2.7207 | -94.725
2.7244 | -94.7104
2.7282 | -94.6281
2.7319 | -94.5235
2.7356 | -94.4453
2.7393 | -94.4293
2.7431 | -94.4687
2.7468 | -94.5413
2.7505 | -94.6263
2.7543 | -94.6904
2.758 | -94.7219
2.7617 | -94.7205
2.7654 | -94.6667
2.7692 | -94.574
2.7729 | -94.4683
2.7766 | -94.3728
2.7803 | -94.3272
2.7841 | -94.3279
2.7878 | -94.3483
2.7915 | -94.3658
2.7952 | -94.3692
2.799 | -94.3558
2.8027 | -94.3566
2.8064 | -94.3645
2.8102 | -94.3662
2.8139 | -94.3476
2.8176 | -94.29
2.8213 | -94.2076
2.8251 | -94.1417
2.8288 | -94.1175
2.8325 | -94.1564
2.8362 | -94.2697
2.84 | -94.4229
2.8437 | -94.5837
2.8474 | -94.7291
2.8512 | -94.8264
2.8549 | -94.8721
2.8586 | -94.8706
2.8623 | -94.8368
2.8661 | -94.7873
2.8698 | -94.7333
2.8735 | -94.6655
2.8772 | -94.5861
2.881 | -94.5048
2.8847 | -94.4283
2.8884 | -94.3915
2.8921 | -94.4055
2.8959 | -94.4693
2.8996 | -94.5757
2.9033 | -94.6959
2.9071 | -94.7999
2.9108 | -94.8604
2.9145 | -94.8673
2.9182 | -94.8188
2.922 | -94.7468
2.9257 | -94.6935
2.9294 | -94.6742
2.9331 | -94.6834
2.9369 | -94.7152
2.9406 | -94.7534
2.9443 | -94.7882
2.9481 | -94.8323
2.9518 | -94.8947
2.9555 | -94.9827
2.9592 | -95.095
2.963 | -95.2198
2.9667 | -95.3206
2.9704 | -95.3858
2.9741 | -95.3954
2.9779 | -95.3679
2.9816 | -95.349
2.9853 | -95.3495
2.9891 | -95.3978
2.9928 | -95.4712
2.9965 | -95.5248
3.0002 | -95.52
3.004 | -95.4538
3.0077 | -95.3546
3.0114 | -95.2532
3.0151 | -95.1815
3.0189 | -95.1292
3.0226 | -95.1048
3.0263 | -95.0939
3.03 | -95.06
3.0338 | -95.002
3.0375 | -94.909
3.0412 | -94.8079
3.045 | -94.7227
3.0487 | -94.6767
3.0524 | -94.6828
3.0561 | -94.7559
3.0599 | -94.8934
3.0636 | -95.0564
3.0673 | -95.1989
3.071 | -95.2736
3.0748 | -95.2562
3.0785 | -95.1539
3.0822 | -95.0066
3.086 | -94.8756
3.0897 | -94.8111
3.0934 | -94.8303
3.0971 | -94.9192
3.1009 | -95.0385
3.1046 | -95.1334
3.1083 | -95.1989
3.112 | -95.2566
3.1158 | -95.3218
3.1195 | -95.4072
3.1232 | -95.5049
3.127 | -95.5919
3.1307 | -95.6452
3.1344 | -95.6692
3.1381 | -95.6771
3.1419 | -95.6784
3.1456 | -95.664
3.1493 | -95.6235
3.153 | -95.5897
3.1568 | -95.5588
3.1605 | -95.5409
3.1642 | -95.5371
3.1679 | -95.5394
3.1717 | -95.5683
3.1754 | -95.626
3.1791 | -95.7046
3.1829 | -95.7948
3.1866 | -95.8648
3.1903 | -95.8868
3.194 | -95.8508
3.1978 | -95.7447
3.2015 | -95.5612
3.2052 | -95.3448
3.2089 | -95.1573
3.2127 | -95.0421
3.2164 | -95.028
3.2201 | -95.1145
3.2239 | -95.3594
3.2276 | -95.5713
3.2313 | -95.8145
3.235 | -96.0551
3.2388 | -96.2499
3.2425 | -96.3556
3.2462 | -96.3481
3.2499 | -96.2393
3.2537 | -96.0671
3.2574 | -95.8758
3.2611 | -95.7035
3.2649 | -95.5771
3.2686 | -95.5109
3.2723 | -95.4944
3.276 | -95.5175
3.2798 | -95.5709
3.2835 | -95.6431
3.2872 | -95.7253
3.2909 | -95.812
3.2947 | -95.9059
3.2984 | -95.9872
3.3021 | -96.0466
3.3058 | -96.0956
3.3096 | -96.1241
3.3133 | -96.1289
3.317 | -96.1029
3.3208 | -96.0411
3.3245 | -95.9382
3.3282 | -95.8289
3.3319 | -95.7257
3.3357 | -95.6467
3.3394 | -95.6107
3.3431 | -95.629
3.3468 | -95.6871
3.3506 | -95.7698
3.3543 | -95.8825
3.358 | -95.9869
3.3618 | -96.0487
3.3655 | -96.0645
3.3692 | -96.0273
3.3729 | -95.9506
3.3767 | -95.8745
3.3804 | -95.817
3.3841 | -95.7843
3.3878 | -95.7868
3.3916 | -95.8184
3.3953 | -95.8728
3.399 | -95.9377
3.4027 | -96.0119
3.4065 | -96.0843
3.4102 | -96.1351
3.4139 | -96.1663
3.4177 | -96.1889
3.4214 | -96.198
3.4251 | -96.1931
3.4288 | -96.1875
3.4326 | -96.1761
3.4363 | -96.1628
3.44 | -96.136
3.4437 | -96.1087
3.4475 | -96.0874
3.4512 | -96.0832
3.4549 | -96.1011
3.4587 | -96.1305
3.4624 | -96.1695
3.4661 | -96.2101
3.4698 | -96.2591
3.4736 | -96.3002
3.4773 | -96.3379
3.481 | -96.3728
3.4847 | -96.3894
3.4885 | -96.4003
3.4922 | -96.3946
3.4959 | -96.376
3.4997 | -96.3417
3.5034 | -96.2991
3.5071 | -96.2533
3.5108 | -96.2214
3.5146 | -96.2227
3.5183 | -96.2552
3.522 | -96.3184
3.5257 | -96.3962
3.5295 | -96.4809
3.5332 | -96.5555
3.5369 | -96.626
3.5406 | -96.7071
3.5444 | -96.7943
3.5481 | -96.8974
3.5518 | -96.9949
3.5556 | -97.0508
3.5593 | -97.0501
3.563 | -96.9855
3.5667 | -96.8774
3.5705 | -96.7355
3.5742 | -96.5824
3.5779 | -96.4356
3.5816 | -96.3077
3.5854 | -96.2061
3.5891 | -96.1269
3.5928 | -96.0674
3.5966 | -96.0213
3.6003 | -95.9909
3.604 | -95.9788
3.6077 | -95.9824
3.6115 | -96.0013
3.6152 | -96.033
3.6189 | -96.0854
3.6226 | -96.1546
3.6264 | -96.2592
3.6301 | -96.3922
3.6338 | -96.5408
3.6375 | -96.674
3.6413 | -96.7722
3.645 | -96.8169
3.6487 | -96.8031
3.6525 | -96.7574
3.6562 | -96.7049
3.6599 | -96.6547
3.6636 | -96.6101
3.6674 | -96.5594
3.6711 | -96.4933
3.6748 | -96.4228
3.6785 | -96.356
3.6823 | -96.3089
3.686 | -96.2923
3.6897 | -96.3098
3.6935 | -96.3662
3.6972 | -96.4499
3.7009 | -96.5541
3.7046 | -96.6725
3.7084 | -96.783
3.7121 | -96.8621
3.7158 | -96.9146
3.7195 | -96.9472
3.7233 | -96.9657
3.727 | -96.986
3.7307 | -96.9979
3.7345 | -97.0007
3.7382 | -96.9963
3.7419 | -96.9579
3.7456 | -96.8788
3.7494 | -96.7669
3.7531 | -96.6445
3.7568 | -96.552
3.7605 | -96.5886
3.7643 | -96.6086
3.768 | -96.6802
3.7717 | -96.7901
3.7754 | -96.8962
3.7792 | -96.9934
3.7829 | -97.0611
3.7866 | -97.0793
3.7904 | -97.0527
3.7941 | -97.0014
3.7978 | -96.9329
3.8015 | -96.8859
3.8053 | -96.8681
3.809 | -96.8821
3.8127 | -96.9191
3.8164 | -96.9529
3.8202 | -96.9711
3.8239 | -96.974
3.8276 | -96.9621
3.8314 | -96.9469
3.8351 | -96.9469
3.8388 | -96.9734
3.8425 | -97.0146
3.8463 | -97.071
3.85 | -97.1309
3.8537 | -97.1979
3.8574 | -97.2462
3.8612 | -97.2661
3.8649 | -97.2462
3.8686 | -97.196
3.8723 | -97.1285
3.8761 | -97.0596
3.8798 | -97.006
3.8835 | -96.9713
3.8873 | -96.9549
3.891 | -96.9594
3.8947 | -96.9829
3.8984 | -97.0277
3.9022 | -97.0877
3.9059 | -97.1537
3.9096 | -97.2145
3.9133 | -97.2669
3.9171 | -97.296
3.9208 | -97.3048
3.9245 | -97.297
3.9283 | -97.2812
3.932 | -97.2448
3.9357 | -97.2086
3.9394 | -97.1671
3.9432 | -97.1318
3.9469 | -97.0986
3.9506 | -97.066
3.9543 | -97.0319
3.9581 | -97.0029
3.9618 | -96.9803
3.9655 | -96.9755
3.9693 | -96.995
3.973 | -97.0478
3.9767 | -97.1181
3.9804 | -97.192
3.9842 | -97.2753
3.9879 | -97.355
3.9916 | -97.4234
3.9953 | -97.4815
3.9991 | -97.5063
4.0028 | -97.511
4.0065 | -97.4842
4.0102 | -97.4348
4.014 | -97.3698
4.0177 | -97.3119
4.0214 | -97.2726
4.0252 | -97.2595
4.0289 | -97.2715
4.0326 | -97.2939
4.0363 | -97.3154
4.0401 | -97.337
4.0438 | -97.3648
4.0475 | -97.3859
4.0512 | -97.4107
4.055 | -97.4349
4.0587 | -97.4562
4.0624 | -97.4719
4.0662 | -97.4757
4.0699 | -97.4763
4.0736 | -97.4915
4.0773 | -97.5218
4.0811 | -97.5625
4.0848 | -97.6039
4.0885 | -97.6374
4.0922 | -97.6629
4.096 | -97.6833
4.0997 | -97.6925
4.1034 | -97.6804
4.1072 | -97.6578
4.1109 | -97.6291
4.1146 | -97.5925
4.1183 | -97.5675
4.1221 | -97.5542
4.1258 | -97.5569
4.1295 | -97.5833
4.1332 | -97.6296
4.137 | -97.667
4.1407 | -97.6956
4.1444 | -97.7233
4.1481 | -97.7409
4.1519 | -97.7465
4.1556 | -97.7414
4.1593 | -97.7286
4.1631 | -97.7164
4.1668 | -97.7066
4.1705 | -97.6795
4.1742 | -97.6484
4.178 | -97.6168
4.1817 | -97.5889
4.1854 | -97.582
4.1891 | -97.6001
4.1929 | -97.6447
4.1966 | -97.7221
4.2003 | -97.8358
4.2041 | -97.9695
4.2078 | -98.0998
4.2115 | -98.2103
4.2152 | -98.2889
4.219 | -98.3317
4.2227 | -98.3424
4.2264 | -98.3224
4.2301 | -98.2857
4.2339 | -98.2494
4.2376 | -98.2165
4.2413 | -98.1947
4.245 | -98.183
4.2488 | -98.171
4.2525 | -98.1672
4.2562 | -98.1719
4.26 | -98.1704
4.2637 | -98.1595
4.2674 | -98.1427
4.2711 | -98.1155
4.2749 | -98.0829
4.2786 | -98.0295
4.2823 | -97.9627
4.286 | -97.8879
4.2898 | -97.8137
4.2935 | -97.7354
4.2972 | -97.6714
4.301 | -97.6554
4.3047 | -97.6281
4.3084 | -97.615
4.3121 | -97.6191
4.3159 | -97.6405
4.3196 | -97.6681
4.3233 | -97.6959
4.327 | -97.7283
4.3308 | -97.7644
4.3345 | -97.7925
4.3382 | -97.8216
4.342 | -97.857
4.3457 | -97.8889
4.3494 | -97.9189
4.3531 | -97.9525
4.3569 | -97.9803
4.3606 | -98.0003
4.3643 | -97.9994
4.368 | -97.9842
4.3718 | -97.964
4.3755 | -97.9413
4.3792 | -97.9236
4.3829 | -97.9184
4.3867 | -97.9189
4.3904 | -97.9373
4.3941 | -97.9649
4.3979 | -97.999
4.4016 | -98.0342
4.4053 | -98.0785
4.409 | -98.1212
4.4128 | -98.1637
4.4165 | -98.2216
4.4202 | -98.2742
4.4239 | -98.3094
4.4277 | -98.346
4.4314 | -98.3723
4.4351 | -98.3821
4.4389 | -98.3736
4.4426 | -98.3541
4.4463 | -98.3291
4.45 | -98.3003
4.4538 | -98.2834
4.4575 | -98.2928
4.4612 | -98.3175
4.4649 | -98.3562
4.4687 | -98.3973
4.4724 | -98.4394
4.4761 | -98.4712
4.4799 | -98.4829
4.4836 | -98.4781
4.4873 | -98.4619
4.491 | -98.4372
4.4948 | -98.4085
4.4985 | -98.3735
4.5022 | -98.3514
4.5059 | -98.3441
4.5097 | -98.3556
4.5134 | -98.3775
4.5171 | -98.4
4.5208 | -98.4194
4.5246 | -98.4314
4.5283 | -98.4326
4.532 | -98.4254
4.5358 | -98.4217
4.5395 | -98.4209
4.5432 | -98.4252
4.5469 | -98.4386
4.5507 | -98.4697
4.5544 | -98.5025
4.5581 | -98.5352
4.5618 | -98.5725
4.5656 | -98.6089
4.5693 | -98.6356
4.573 | -98.6503
4.5768 | -98.6592
4.5805 | -98.664
4.5842 | -98.6633
4.5879 | -98.6675
4.5917 | -98.6748
4.5954 | -98.6669
4.5991 | -98.6524
4.6028 | -98.6415
4.6066 | -98.6225
4.6103 | -98.6055
4.614 | -98.5895
4.6178 | -98.5764
4.6215 | -98.571
4.6252 | -98.5774
4.6289 | -98.6005
4.6327 | -98.6364
4.6364 | -98.6776
4.6401 | -98.7142
4.6438 | -98.7418
4.6476 | -98.7496
4.6513 | -98.7269
4.655 | -98.6843
4.6587 | -98.6258
4.6625 | -98.5693
4.6662 | -98.5375
4.6699 | -98.5336
4.6737 | -98.5542
4.6774 | -98.6043
4.6811 | -98.6758
4.6848 | -98.7452
4.6886 | -98.797
4.6923 | -98.8204
4.696 | -98.8039
4.6997 | -98.7492
4.7035 | -98.6734
4.7072 | -98.5952
4.7109 | -98.5298
4.7147 | -98.4882
4.7184 | -98.4827
4.7221 | -98.5193
4.7258 | -98.5899
4.7296 | -98.6843
4.7333 | -98.7968
4.737 | -98.9081
4.7407 | -99.005
4.7445 | -99.0868
4.7482 | -99.1459
4.7519 | -99.1699
4.7556 | -99.1826
4.7594 | -99.183
4.7631 | -99.1759
4.7668 | -99.1653
4.7706 | -99.1469
4.7743 | -99.1221
4.778 | -99.0931
4.7817 | -99.0633
4.7855 | -99.0351
4.7892 | -98.9998
4.7929 | -98.9624
4.7966 | -98.919
4.8004 | -98.8725
4.8041 | -98.8238
4.8078 | -98.7867
4.8116 | -98.7667
4.8153 | -98.7645
4.819 | -98.7929
4.8227 | -98.8539
4.8265 | -98.9284
4.8302 | -99.0127
4.8339 | -99.102
4.8376 | -99.172
4.8414 | -99.2372
4.8451 | -99.2909
4.8488 | -99.3299
4.8526 | -99.3531
4.8563 | -99.3637
4.86 | -99.3635
4.8637 | -99.3594
4.8675 | -99.3461
4.8712 | -99.3293
4.8749 | -99.3122
4.8786 | -99.294
4.8824 | -99.269
4.8861 | -99.2412
4.8898 | -99.2051
4.8935 | -99.1641
4.8973 | -99.1252
4.901 | -99.0998
4.9047 | -99.0828
4.9085 | -99.0611
4.9122 | -99.0422
4.9159 | -99.0339
4.9196 | -99.026
4.9234 | -99.0192
4.9271 | -99.0072
4.9308 | -98.996
4.9345 | -98.9861
4.9383 | -98.9745
4.942 | -98.9611
4.9457 | -98.952
4.9495 | -98.9387
4.9532 | -98.9199
4.9569 | -98.8956
4.9606 | -98.8755
4.9644 | -98.8563
4.9681 | -98.8402
4.9718 | -98.821
4.9755 | -98.8033
4.9793 | -98.7948
4.983 | -98.7925
4.9867 | -98.7988
4.9905 | -98.8178
4.9942 | -98.8294
4.9979 | -98.8361
5.0016 | -98.8393
5.0054 | -98.8306
5.0091 | -98.8056
5.0128 | -98.7807
5.0165 | -98.7602
5.0203 | -98.7442
5.024 | -98.7392
5.0277 | -98.7521
5.0314 | -98.7862
5.0352 | -98.8424
5.0389 | -98.9107
5.0426 | -98.9885
5.0464 | -99.063
5.0501 | -99.1326
5.0538 | -99.1896
5.0575 | -99.2301
5.0613 | -99.256
5.065 | -99.2692
5.0687 | -99.2515
5.0724 | -99.2212
5.0762 | -99.1839
5.0799 | -99.1308
5.0836 | -99.0822
5.0874 | -99.0405
5.0911 | -99.0019
5.0948 | -98.978
5.0985 | -98.9773
5.1023 | -98.9981
5.106 | -99.0455
5.1097 | -99.1097
5.1134 | -99.1889
5.1172 | -99.2697
5.1209 | -99.3381
5.1246 | -99.3969
5.1283 | -99.4276
5.1321 | -99.4151
5.1358 | -99.3725
5.1395 | -99.3141
5.1433 | -99.2378
5.147 | -99.1646
5.1507 | -99.0986
5.1544 | -99.0494
5.1582 | -99.0302
5.1619 | -99.0352
5.1656 | -99.0692
5.1693 | -99.1379
5.1731 | -99.2245
5.1768 | -99.3153
5.1805 | -99.4037
5.1843 | -99.4846
5.188 | -99.5539
5.1917 | -99.6076
5.1954 | -99.6554
5.1992 | -99.6927
5.2029 | -99.7146
5.2066 | -99.7252
5.2103 | -99.7325
5.2141 | -99.736
5.2178 | -99.7297
5.2215 | -99.7196
5.2252 | -99.7073
5.229 | -99.6914
5.2327 | -99.6633
5.2364 | -99.6321
5.2402 | -99.602
5.2439 | -99.5626
5.2476 | -99.507
5.2513 | -99.4366
5.2551 | -99.3587
5.2588 | -99.2797
5.2625 | -99.2042
5.2662 | -99.1425
5.27 | -99.104
5.2737 | -99.0821
5.2774 | -99.089
5.2812 | -99.1111
5.2849 | -99.1466
5.2886 | -99.1923
5.2923 | -99.2304
5.2961 | -99.2652
5.2998 | -99.2924
5.3035 | -99.3068
5.3072 | -99.3167
5.311 | -99.3225
5.3147 | -99.3241
5.3184 | -99.3246
5.3222 | -99.321
5.3259 | -99.3216
5.3296 | -99.3331
5.3333 | -99.3471
5.3371 | -99.3561
5.3408 | -99.3694
5.3445 | -99.3898
5.3482 | -99.4104
5.352 | -99.4401
5.3557 | -99.4658
5.3594 | -99.4975
5.3631 | -99.5336
5.3669 | -99.5704
5.3706 | -99.6258
5.3743 | -99.7125
5.3781 | -99.7908
5.3818 | -99.8744
5.3855 | -99.959
5.3892 | -100.0358
5.393 | -100.1014
5.3967 | -100.1529
5.4004 | -100.1877
5.4041 | -100.2064
5.4079 | -100.2021
5.4116 | -100.1754
5.4153 | -100.1331
5.4191 | -100.092
5.4228 | -100.0473
5.4265 | -100.001
5.4302 | -99.9679
5.434 | -99.9426
5.4377 | -99.9333
5.4414 | -99.9386
5.4451 | -99.9564
5.4489 | -99.983
5.4526 | -100.0169
5.4563 | -100.0559
5.4601 | -100.0868
5.4638 | -100.1121
5.4675 | -100.1345
5.4712 | -100.1373
5.475 | -100.1245
5.4787 | -100.0964
5.4824 | -100.0531
5.4861 | -100.003
5.4899 | -99.9529
5.4936 | -99.9006
5.4973 | -99.8524
5.501 | -99.8108
5.5048 | -99.7763
5.5085 | -99.7524
5.5122 | -99.7437
5.516 | -99.7488
5.5197 | -99.7734
5.5234 | -99.8113
5.5271 | -99.8582
5.5309 | -99.9071
5.5346 | -99.957
5.5383 | -100.002
5.542 | -100.0379
5.5458 | -100.0632
5.5495 | -100.0757
5.5532 | -100.0677
5.557 | -100.0551
5.5607 | -100.032
5.5644 | -99.9985
5.5681 | -99.9569
5.5719 | -99.9116
5.5756 | -99.8656
5.5793 | -99.832
5.583 | -99.8066
5.5868 | -99.7901
5.5905 | -99.786
5.5942 | -99.7934
5.5979 | -99.8058
5.6017 | -99.8189
5.6054 | -99.8344
5.6091 | -99.8577
5.6129 | -99.8845
5.6166 | -99.9171
5.6203 | -99.945
5.624 | -99.9783
5.6278 | -100.0184
5.6315 | -100.0653
5.6352 | -100.1147
5.6389 | -100.1625
5.6427 | -100.2074
5.6464 | -100.2492
5.6501 | -100.2915
5.6539 | -100.3273
5.6576 | -100.3504
5.6613 | -100.3621
5.665 | -100.3529
5.6688 | -100.3287
5.6725 | -100.2921
5.6762 | -100.238
5.6799 | -100.1805
5.6837 | -100.1227
5.6874 | -100.0613
5.6911 | -99.9991
5.6949 | -99.9476
5.6986 | -99.9132
5.7023 | -99.8871
5.706 | -99.8739
5.7098 | -99.8596
5.7135 | -99.8611
5.7172 | -99.8751
5.7209 | -99.9013
5.7247 | -99.9412
5.7284 | -99.9908
5.7321 | -100.0487
5.7358 | -100.1074
5.7396 | -100.1679
5.7433 | -100.2305
5.747 | -100.2832
5.7508 | -100.3235
5.7545 | -100.3529
5.7582 | -100.3746
5.7619 | -100.3897
5.7657 | -100.3955
5.7694 | -100.3941
5.7731 | -100.3987
5.7768 | -100.3991
5.7806 | -100.3831
5.7843 | -100.3588
5.788 | -100.3335
5.7918 | -100.3029
5.7955 | -100.2652
5.7992 | -100.2136
5.8029 | -100.1694
5.8067 | -100.1356
5.8104 | -100.1066
5.8141 | -100.0752
5.8178 | -100.0431
5.8216 | -100.0135
5.8253 | -99.9831
5.829 | -99.9467
5.8328 | -99.9103
5.8365 | -99.8775
5.8402 | -99.843
5.8439 | -99.807
5.8477 | -99.78
5.8514 | -99.762
5.8551 | -99.7434
5.8588 | -99.727
5.8626 | -99.7245
5.8663 | -99.7322
5.87 | -99.7447
5.8737 | -99.7569
5.8775 | -99.7621
5.8812 | -99.7656
5.8849 | -99.7593
5.8887 | -99.7357
5.8924 | -99.7047
5.8961 | -99.6818
5.8998 | -99.6754
5.9036 | -99.6856
5.9073 | -99.7124
5.911 | -99.7912
5.9147 | -99.8407
5.9185 | -99.9023
5.9222 | -99.9722
5.9259 | -100.0409
5.9297 | -100.1088
5.9334 | -100.1709
5.9371 | -100.2207
5.9408 | -100.2566
5.9446 | -100.2817
5.9483 | -100.286
5.952 | -100.2769
5.9557 | -100.2666
5.9595 | -100.2467
5.9632 | -100.2234
5.9669 | -100.2009
5.9707 | -100.1777
5.9744 | -100.159
5.9781 | -100.1453
5.9818 | -100.138
5.9856 | -100.1365
5.9893 | -100.1298
5.993 | -100.1257
5.9967 | -100.1267
6.0005 | -100.1318
6.0042 | -100.1392
6.0079 | -100.1467
6.0116 | -100.1524
6.0154 | -100.1493
6.0191 | -100.1533
6.0228 | -100.1568
6.0266 | -100.1635
6.0303 | -100.1763
6.034 | -100.1867
6.0377 | -100.1937
6.0415 | -100.2043
6.0452 | -100.1974
6.0489 | -100.1881
6.0526 | -100.1911
6.0564 | -100.1978
6.0601 | -100.2054
6.0638 | -100.21
6.0676 | -100.2217
6.0713 | -100.2345
6.075 | -100.2471
6.0787 | -100.2609
6.0825 | -100.2838
6.0862 | -100.3059
6.0899 | -100.3249
6.0936 | -100.3439
6.0974 | -100.3616
6.1011 | -100.3814
6.1048 | -100.4048
6.1085 | -100.4331
6.1123 | -100.4593
6.116 | -100.4909
6.1197 | -100.5242
6.1235 | -100.5539
6.1272 | -100.5876
6.1309 | -100.6201
6.1346 | -100.6368
6.1384 | -100.6527
6.1421 | -100.6505
6.1458 | -100.6349
6.1495 | -100.6184
6.1533 | -100.604
6.157 | -100.5937
6.1607 | -100.5822
6.1645 | -100.5759
6.1682 | -100.5706
6.1719 | -100.5637
6.1756 | -100.5656
6.1794 | -100.572
6.1831 | -100.5816
6.1868 | -100.5941
6.1905 | -100.6111
6.1943 | -100.631
6.198 | -100.6437
6.2017 | -100.6507
6.2055 | -100.6553
6.2092 | -100.6526
6.2129 | -100.642
6.2166 | -100.6221
6.2204 | -100.6007
6.2241 | -100.5752
6.2278 | -100.5481
6.2315 | -100.5308
6.2353 | -100.5206
6.239 | -100.513
6.2427 | -100.5067
6.2464 | -100.5039
6.2502 | -100.5131
6.2539 | -100.5277
6.2576 | -100.5468
6.2614 | -100.5747
6.2651 | -100.602
6.2688 | -100.623
6.2725 | -100.6503
6.2763 | -100.6797
6.28 | -100.7094
6.2837 | -100.7378
6.2874 | -100.7652
6.2912 | -100.7916
6.2949 | -100.8194
6.2986 | -100.8409
6.3024 | -100.8648
6.3061 | -100.8859
6.3098 | -100.906
6.3135 | -100.9168
6.3173 | -100.9279
6.321 | -100.9339
6.3247 | -100.9368
6.3284 | -100.9464
6.3322 | -100.9578
6.3359 | -100.9729
6.3396 | -100.9951
6.3434 | -101.0144
6.3471 | -101.0455
6.3508 | -101.0866
6.3545 | -101.134
6.3583 | -101.1882
6.362 | -101.2398
6.3657 | -101.2848
6.3694 | -101.3165
6.3732 | -101.3251
6.3769 | -101.3218
6.3806 | -101.2998
6.3843 | -101.2572
6.3881 | -101.2016
6.3918 | -101.1402
6.3955 | -101.0753
6.3993 | -101.0128
6.403 | -100.9543
6.4067 | -100.9014
6.4104 | -100.862
6.4142 | -100.8377
6.4179 | -100.8162
6.4216 | -100.8088
6.4253 | -100.8211
6.4291 | -100.8424
6.4328 | -100.8688
6.4365 | -100.8971
6.4403 | -100.9186
6.444 | -100.9384
6.4477 | -100.951
6.4514 | -100.9445
6.4552 | -100.9486
6.4589 | -100.9464
6.4626 | -100.939
6.4663 | -100.9278
6.4701 | -100.9218
6.4738 | -100.9153
6.4775 | -100.9096
6.4812 | -100.8983
6.485 | -100.8813
6.4887 | -100.8632
6.4924 | -100.8401
6.4962 | -100.8161
6.4999 | -100.7854
6.5036 | -100.7535
6.5073 | -100.718
6.5111 | -100.6858
6.5148 | -100.6597
6.5185 | -100.6369
6.5222 | -100.6219
6.526 | -100.6199
6.5297 | -100.629
6.5334 | -100.6503
6.5372 | -100.6847
6.5409 | -100.731
6.5446 | -100.7855
6.5483 | -100.8421
6.5521 | -100.9034
6.5558 | -100.9573
6.5595 | -101.0156
6.5632 | -101.0708
6.567 | -101.1213
6.5707 | -101.1558
6.5744 | -101.1844
6.5781 | -101.2116
6.5819 | -101.2427
6.5856 | -101.2637
6.5893 | -101.277
6.5931 | -101.2926
6.5968 | -101.3053
6.6005 | -101.3165
6.6042 | -101.3197
6.608 | -101.314
6.6117 | -101.3086
6.6154 | -101.2944
6.6191 | -101.276
6.6229 | -101.2569
6.6266 | -101.2416
6.6303 | -101.2242
6.6341 | -101.2047
6.6378 | -101.1791
6.6415 | -101.1453
6.6452 | -101.115
6.649 | -101.0858
6.6527 | -101.0566
6.6564 | -101.033
6.6601 | -101.0047
6.6639 | -100.9796
6.6676 | -100.9663
6.6713 | -100.9633
6.6751 | -100.9594
6.6788 | -100.9652
6.6825 | -100.9761
6.6862 | -100.9957
6.69 | -101.0239
6.6937 | -101.0547
6.6974 | -101.088
6.7011 | -101.1298
6.7049 | -101.1735
6.7086 | -101.2203
6.7123 | -101.2663
6.716 | -101.3207
6.7198 | -101.3692
6.7235 | -101.4012
6.7272 | -101.4223
6.731 | -101.4418
6.7347 | -101.4492
6.7384 | -101.4471
6.7421 | -101.4396
6.7459 | -101.4252
6.7496 | -101.4066
6.7533 | -101.3867
6.757 | -101.3674
6.7608 | -101.3534
6.7645 | -101.3394
6.7682 | -101.3255
6.772 | -101.3217
6.7757 | -101.3234
6.7794 | -101.3308
6.7831 | -101.3484
6.7869 | -101.3761
6.7906 | -101.4125
6.7943 | -101.4552
6.798 | -101.5023
6.8018 | -101.5482
6.8055 | -101.5954
6.8092 | -101.6317
6.813 | -101.6583
6.8167 | -101.669
6.8204 | -101.6669
6.8241 | -101.6535
6.8279 | -101.6186
6.8316 | -101.5637
6.8353 | -101.5061
6.839 | -101.4386
6.8428 | -101.3684
6.8465 | -101.2975
6.8502 | -101.232
6.8539 | -101.1783
6.8577 | -101.139
6.8614 | -101.1149
6.8651 | -101.1108
6.8689 | -101.1211
6.8726 | -101.1382
6.8763 | -101.1659
6.88 | -101.1975
6.8838 | -101.2323
6.8875 | -101.2606
6.8912 | -101.2831
6.8949 | -101.3077
6.8987 | -101.3246
6.9024 | -101.3304
6.9061 | -101.3314
6.9099 | -101.3283
6.9136 | -101.3145
6.9173 | -101.3003
6.921 | -101.288
6.9248 | -101.2716
6.9285 | -101.254
6.9322 | -101.2436
6.9359 | -101.2343
6.9397 | -101.2327
6.9434 | -101.2406
6.9471 | -101.2607
6.9508 | -101.2888
6.9546 | -101.3172
6.9583 | -101.3492
6.962 | -101.3855
6.9658 | -101.4209
6.9695 | -101.4537
6.9732 | -101.4835
6.9769 | -101.5112
6.9807 | -101.5299
6.9844 | -101.5378
6.9881 | -101.5208
6.9918 | -101.5155
6.9956 | -101.5061
6.9993 | -101.4954
7.003 | -101.4851
7.0068 | -101.47
7.0105 | -101.4543
7.0142 | -101.4421
7.0179 | -101.4336
7.0217 | -101.4239
7.0254 | -101.4172
7.0291 | -101.413
7.0328 | -101.4139
7.0366 | -101.4171
7.0403 | -101.4208
7.044 | -101.4256
7.0478 | -101.4336
7.0515 | -101.4474
7.0552 | -101.4589
7.0589 | -101.4793
7.0627 | -101.5045
7.0664 | -101.527
7.0701 | -101.5487
7.0738 | -101.5695
7.0776 | -101.5845
7.0813 | -101.5903
7.085 | -101.5945
7.0887 | -101.6018
7.0925 | -101.6034
7.0962 | -101.6
7.0999 | -101.5952
7.1037 | -101.586
7.1074 | -101.5792
7.1111 | -101.5759
7.1148 | -101.582
7.1186 | -101.5879
7.1223 | -101.5945
7.126 | -101.6079
7.1297 | -101.6237
7.1335 | -101.6383
7.1372 | -101.6567
7.1409 | -101.6711
7.1447 | -101.6866
7.1484 | -101.7033
7.1521 | -101.719
7.1558 | -101.7299
7.1596 | -101.7381
7.1633 | -101.7493
7.167 | -101.7476
7.1707 | -101.7451
7.1745 | -101.7484
7.1782 | -101.752
7.1819 | -101.7564
7.1857 | -101.7661
7.1894 | -101.7816
7.1931 | -101.7916
7.1968 | -101.7996
7.2006 | -101.8076
7.2043 | -101.8106
7.208 | -101.8057
7.2117 | -101.7928
7.2155 | -101.7745
7.2192 | -101.7522
7.2229 | -101.7215
7.2266 | -101.693
7.2304 | -101.6667
7.2341 | -101.64
7.2378 | -101.6186
7.2416 | -101.6042
7.2453 | -101.5983
7.249 | -101.6037
7.2527 | -101.6085
7.2565 | -101.6245
7.2602 | -101.659
7.2639 | -101.7003
7.2676 | -101.7453
7.2714 | -101.7973
7.2751 | -101.8526
7.2788 | -101.9014
7.2826 | -101.9406
7.2863 | -101.9758
7.29 | -102.0017
7.2937 | -102.013
7.2975 | -102.0125
7.3012 | -102.0058
7.3049 | -101.9934
7.3086 | -101.9752
7.3124 | -101.9568
7.3161 | -101.9383
7.3198 | -101.925
7.3236 | -101.9133
7.3273 | -101.9083
7.331 | -101.9109
7.3347 | -101.9175
7.3385 | -101.9323
7.3422 | -101.9467
7.3459 | -101.9653
7.3496 | -101.9861
7.3534 | -102.0048
7.3571 | -102.0183
7.3608 | -102.028
7.3645 | -102.0274
7.3683 | -102.0146
7.372 | -102.0007
7.3757 | -101.986
7.3795 | -101.9704
7.3832 | -101.9492
7.3869 | -101.9285
7.3906 | -101.9154
7.3944 | -101.9017
7.3981 | -101.8869
7.4018 | -101.876
7.4055 | -101.872
7.4093 | -101.8762
7.413 | -101.8783
7.4167 | -101.8835
7.4205 | -101.8987
7.4242 | -101.9212
7.4279 | -101.9451
7.4316 | -101.9714
7.4354 | -102.0008
7.4391 | -102.0327
7.4428 | -102.0612
7.4465 | -102.082
7.4503 | -102.0995
7.454 | -102.1075
7.4577 | -102.1084
7.4614 | -102.1047
7.4652 | -102.0893
7.4689 | -102.0684
7.4726 | -102.0412
7.4764 | -102.0025
7.4801 | -101.961
7.4838 | -101.9177
7.4875 | -101.8668
7.4913 | -101.8244
7.495 | -101.7889
7.4987 | -101.7627
7.5024 | -101.7466
7.5062 | -101.7433
7.5099 | -101.7473
7.5136 | -101.759
7.5174 | -101.7819
7.5211 | -101.8166
7.5248 | -101.855
7.5285 | -101.8914
7.5323 | -101.9231
7.536 | -101.9529
7.5397 | -101.9802
7.5434 | -102.0064
7.5472 | -102.0255
7.5509 | -102.0388
7.5546 | -102.0453
7.5584 | -102.0473
7.5621 | -102.045
7.5658 | -102.0445
7.5695 | -102.0409
7.5733 | -102.0364
7.577 | -102.0371
7.5807 | -102.0417
7.5844 | -102.0512
7.5882 | -102.0651
7.5919 | -102.0802
7.5956 | -102.099
7.5993 | -102.1276
7.6031 | -102.16
7.6068 | -102.1945
7.6105 | -102.2262
7.6143 | -102.2591
7.618 | -102.2936
7.6217 | -102.3358
7.6254 | -102.3692
7.6292 | -102.3933
7.6329 | -102.4142
7.6366 | -102.4307
7.6403 | -102.4395
7.6441 | -102.4351
7.6478 | -102.4187
7.6515 | -102.4001
7.6553 | -102.3784
7.659 | -102.3504
7.6627 | -102.3215
7.6664 | -102.2914
7.6702 | -102.2675
7.6739 | -102.2488
7.6776 | -102.232
7.6813 | -102.2137
7.6851 | -102.2005
7.6888 | -102.1966
7.6925 | -102.1993
7.6963 | -102.1984
7.7 | -102.1989
7.7037 | -102.1984
7.7074 | -102.1927
7.7112 | -102.1823
7.7149 | -102.1609
7.7186 | -102.1313
7.7223 | -102.0983
7.7261 | -102.0637
7.7298 | -102.0293
7.7335 | -101.9926
7.7372 | -101.951
7.741 | -101.9183
7.7447 | -101.8889
7.7484 | -101.8623
7.7522 | -101.843
7.7559 | -101.8333
7.7596 | -101.8325
7.7633 | -101.8413
7.7671 | -101.8583
7.7708 | -101.8852
7.7745 | -101.9191
7.7782 | -101.9611
7.782 | -102.007
7.7857 | -102.0521
7.7894 | -102.0952
7.7932 | -102.1409
7.7969 | -102.1836
7.8006 | -102.2194
7.8043 | -102.2468
7.8081 | -102.2668
7.8118 | -102.2825
7.8155 | -102.2865
7.8192 | -102.2882
7.823 | -102.2901
7.8267 | -102.2881
7.8304 | -102.2895
7.8341 | -102.291
7.8379 | -102.2859
7.8416 | -102.283
7.8453 | -102.2848
7.8491 | -102.2887
7.8528 | -102.2956
7.8565 | -102.3013
7.8602 | -102.3181
7.864 | -102.3377
7.8677 | -102.3495
7.8714 | -102.3634
7.8751 | -102.3767
7.8789 | -102.3873
7.8826 | -102.3953
7.8863 | -102.4016
7.8901 | -102.4125
7.8938 | -102.4169
7.8975 | -102.4174
7.9012 | -102.4182
7.905 | -102.4144
7.9087 | -102.3981
7.9124 | -102.3768
7.9161 | -102.3498
7.9199 | -102.3172
7.9236 | -102.2829
7.9273 | -102.246
7.9311 | -102.2045
7.9348 | -102.1545
7.9385 | -102.1017
7.9422 | -102.0479
7.946 | -101.9949
7.9497 | -101.9488
7.9534 | -101.9146
7.9571 | -101.8822
7.9609 | -101.852
7.9646 | -101.8348
7.9683 | -101.8296
7.972 | -101.8349
7.9758 | -101.8431
7.9795 | -101.8601
7.9832 | -101.8847
7.987 | -101.9076
7.9907 | -101.9317
7.9944 | -101.9595
7.9981 | -101.9848
8.0019 | -102.0114
8.0056 | -102.0374
8.0093 | -102.0668
8.013 | -102.0952
8.0168 | -102.123
8.0205 | -102.1565
8.0242 | -102.191
8.028 | -102.2249
8.0317 | -102.2556
8.0354 | -102.2802
8.0391 | -102.3009
8.0429 | -102.319
8.0466 | -102.3334
8.0503 | -102.3455
8.054 | -102.351
8.0578 | -102.3531
8.0615 | -102.3493
8.0652 | -102.3422
8.0689 | -102.3411
8.0727 | -102.3427
8.0764 | -102.3491
8.0801 | -102.3616
8.0839 | -102.373
8.0876 | -102.3874
8.0913 | -102.409
8.095 | -102.4337
8.0988 | -102.4603
8.1025 | -102.4888
8.1062 | -102.5127
8.1099 | -102.5343
8.1137 | -102.5485
8.1174 | -102.565
8.1211 | -102.5741
8.1249 | -102.5765
8.1286 | -102.5745
8.1323 | -102.5649
8.136 | -102.5527
8.1398 | -102.5362
8.1435 | -102.517
8.1472 | -102.4918
8.1509 | -102.4656
8.1547 | -102.4422
8.1584 | -102.4233
8.1621 | -102.4104
8.1659 | -102.4002
8.1696 | -102.3989
8.1733 | -102.4049
8.177 | -102.4106
8.1808 | -102.423
8.1845 | -102.44
8.1882 | -102.4571
8.1919 | -102.4724
8.1957 | -102.4844
8.1994 | -102.4941
8.2031 | -102.4998
8.2068 | -102.5079
8.2106 | -102.5112
8.2143 | -102.5083
8.218 | -102.5028
8.2218 | -102.4912
8.2255 | -102.478
8.2292 | -102.4617
8.2329 | -102.4398
8.2367 | -102.4153
8.2404 | -102.3842
8.2441 | -102.3546
8.2478 | -102.3214
8.2516 | -102.2886
8.2553 | -102.2649
8.259 | -102.2388
8.2628 | -102.216
8.2665 | -102.206
8.2702 | -102.2024
8.2739 | -102.2011
8.2777 | -102.2039
8.2814 | -102.2064
8.2851 | -102.2121
8.2888 | -102.2203
8.2926 | -102.224
8.2963 | -102.227
8.3 | -102.2339
8.3038 | -102.2355
8.3075 | -102.2318
8.3112 | -102.2279
8.3149 | -102.2222
8.3187 | -102.218
8.3224 | -102.2121
8.3261 | -102.2028
8.3298 | -102.1973
8.3336 | -102.1937
8.3373 | -102.1908
8.341 | -102.1922
8.3447 | -102.2004
8.3485 | -102.2066
8.3522 | -102.2109
8.3559 | -102.2169
8.3597 | -102.2216
8.3634 | -102.2301
8.3671 | -102.2384
8.3708 | -102.2473
8.3746 | -102.2571
8.3783 | -102.2649
8.382 | -102.2723
8.3857 | -102.2793
8.3895 | -102.2798
8.3932 | -102.2822
8.3969 | -102.2845
8.4007 | -102.2833
8.4044 | -102.2833
8.4081 | -102.285
8.4118 | -102.2918
8.4156 | -102.3003
8.4193 | -102.3087
8.423 | -102.3221
8.4267 | -102.3409
8.4305 | -102.3678
8.4342 | -102.3993
8.4379 | -102.4377
8.4416 | -102.4761
8.4454 | -102.5173
8.4491 | -102.5556
8.4528 | -102.5848
8.4566 | -102.6142
8.4603 | -102.6425
8.464 | -102.6626
8.4677 | -102.6731
8.4715 | -102.6767
8.4752 | -102.6783
8.4789 | -102.6692
8.4826 | -102.65
8.4864 | -102.6248
8.4901 | -102.6006
8.4938 | -102.5775
8.4976 | -102.5527
8.5013 | -102.5247
8.505 | -102.5046
8.5087 | -102.489
8.5125 | -102.4789
8.5162 | -102.4706
8.5199 | -102.4688
8.5236 | -102.4698
8.5274 | -102.4759
8.5311 | -102.4849
8.5348 | -102.4964
8.5386 | -102.5131
8.5423 | -102.5355
8.546 | -102.5552
8.5497 | -102.5774
8.5535 | -102.5989
8.5572 | -102.6229
8.5609 | -102.6434
8.5646 | -102.6669
8.5684 | -102.6903
8.5721 | -102.7146
8.5758 | -102.737
8.5795 | -102.7552
8.5833 | -102.7718
8.587 | -102.7907
8.5907 | -102.8089
8.5945 | -102.8235
8.5982 | -102.831
8.6019 | -102.8332
8.6056 | -102.8394
8.6094 | -102.8456
8.6131 | -102.8476
8.6168 | -102.8448
8.6205 | -102.8403
8.6243 | -102.8348
8.628 | -102.8331
8.6317 | -102.829
8.6355 | -102.8212
8.6392 | -102.8137
8.6429 | -102.8115
8.6466 | -102.8119
8.6504 | -102.8123
8.6541 | -102.8106
8.6578 | -102.8141
8.6615 | -102.8184
8.6653 | -102.8221
8.669 | -102.8299
8.6727 | -102.8452
8.6765 | -102.8618
8.6802 | -102.8797
8.6839 | -102.8961
8.6876 | -102.9174
8.6914 | -102.9408
8.6951 | -102.9625
8.6988 | -102.98
8.7025 | -102.9936
8.7063 | -103.0059
8.71 | -103.0118
8.7137 | -103.0126
8.7174 | -103.0132
8.7212 | -103.0094
8.7249 | -103.0042
8.7286 | -102.989
8.7324 | -102.9745
8.7361 | -102.9572
8.7398 | -102.9371
8.7435 | -102.9179
8.7473 | -102.8995
8.751 | -102.8812
8.7547 | -102.8651
8.7584 | -102.8477
8.7622 | -102.8291
8.7659 | -102.8182
8.7696 | -102.808
8.7734 | -102.7962
8.7771 | -102.7824
8.7808 | -102.7762
8.7845 | -102.7668
8.7883 | -102.7568
8.792 | -102.7522
8.7957 | -102.7538
8.7994 | -102.7534
8.8032 | -102.7558
8.8069 | -102.765
8.8106 | -102.7769
8.8143 | -102.7907
8.8181 | -102.8044
8.8218 | -102.8162
8.8255 | -102.8266
8.8293 | -102.8453
8.833 | -102.8613
8.8367 | -102.869
8.8404 | -102.8772
8.8442 | -102.8823
8.8479 | -102.8831
8.8516 | -102.884
8.8553 | -102.8796
8.8591 | -102.8688
8.8628 | -102.8564
8.8665 | -102.8428
8.8703 | -102.8246
8.874 | -102.8074
8.8777 | -102.7937
8.8814 | -102.7811
8.8852 | -102.7665
8.8889 | -102.7537
8.8926 | -102.7389
8.8963 | -102.7256
8.9001 | -102.7155
8.9038 | -102.7088
8.9075 | -102.7063
8.9113 | -102.7002
8.915 | -102.6978
8.9187 | -102.6935
8.9224 | -102.6935
8.9262 | -102.6941
8.9299 | -102.6996
8.9336 | -102.7091
8.9373 | -102.7218
8.9411 | -102.7395
8.9448 | -102.762
8.9485 | -102.7863
8.9522 | -102.814
8.956 | -102.8454
8.9597 | -102.8692
8.9634 | -102.8929
8.9672 | -102.9186
8.9709 | -102.9425
8.9746 | -102.9636
8.9783 | -102.9802
8.9821 | -102.9952
8.9858 | -103.0095
8.9895 | -103.019
8.9932 | -103.0209
8.997 | -103.0257
9.0007 | -103.0303
9.0044 | -103.025
9.0082 | -103.0111
9.0119 | -102.9974
9.0156 | -102.9855
9.0193 | -102.9749
9.0231 | -102.9673
9.0268 | -102.9539
9.0305 | -102.9399
9.0342 | -102.9278
9.038 | -102.9204
9.0417 | -102.9175
9.0454 | -102.9189
9.0491 | -102.9215
9.0529 | -102.9234
9.0566 | -102.9301
9.0603 | -102.9369
9.0641 | -102.9458
9.0678 | -102.9612
9.0715 | -102.9783
9.0752 | -102.9921
9.079 | -103.0106
9.0827 | -103.0279
9.0864 | -103.0451
9.0901 | -103.0634
9.0939 | -103.0818
9.0976 | -103.1025
9.1013 | -103.12
9.1051 | -103.1348
9.1088 | -103.1512
9.1125 | -103.1646
9.1162 | -103.1751
9.12 | -103.1912
9.1237 | -103.2041
9.1274 | -103.2091
9.1311 | -103.2222
9.1349 | -103.239
9.1386 | -103.2629
9.1423 | -103.2845
9.1461 | -103.3004
9.1498 | -103.3165
9.1535 | -103.3343
9.1572 | -103.3482
9.161 | -103.3602
9.1647 | -103.3739
9.1684 | -103.3841
9.1721 | -103.3903
9.1759 | -103.3938
9.1796 | -103.3954
9.1833 | -103.396
9.187 | -103.3891
9.1908 | -103.378
9.1945 | -103.368
9.1982 | -103.3519
9.202 | -103.3326
9.2057 | -103.3188
9.2094 | -103.3093
9.2131 | -103.2927
9.2169 | -103.2736
9.2206 | -103.2598
9.2243 | -103.248
9.228 | -103.2315
9.2318 | -103.2165
9.2355 | -103.2055
9.2392 | -103.1979
9.243 | -103.1898
9.2467 | -103.1827
9.2504 | -103.1706
9.2541 | -103.1566
9.2579 | -103.1483
9.2616 | -103.1386
9.2653 | -103.1205
9.269 | -103.1036
9.2728 | -103.09
9.2765 | -103.077
9.2802 | -103.0668
9.284 | -103.0614
9.2877 | -103.054
9.2914 | -103.044
9.2951 | -103.0439
9.2989 | -103.0463
9.3026 | -103.0484
9.3063 | -103.0559
9.31 | -103.07
9.3138 | -103.0908
9.3175 | -103.1144
9.3212 | -103.1381
9.3249 | -103.1698
9.3287 | -103.2065
9.3324 | -103.2446
9.3361 | -103.2851
9.3399 | -103.3203
9.3436 | -103.3509
9.3473 | -103.3811
9.351 | -103.4127
9.3548 | -103.4389
9.3585 | -103.4545
9.3622 | -103.4637
9.3659 | -103.4715
9.3697 | -103.4682
9.3734 | -103.4588
9.3771 | -103.4513
9.3809 | -103.4427
9.3846 | -103.4342
9.3883 | -103.4325
9.392 | -103.4313
9.3958 | -103.4273
9.3995 | -103.4299
9.4032 | -103.4395
9.4069 | -103.4516
9.4107 | -103.4653
9.4144 | -103.483
9.4181 | -103.4976
9.4218 | -103.5185
9.4256 | -103.5438
9.4293 | -103.5668
9.433 | -103.5852
9.4368 | -103.6057
9.4405 | -103.6291
9.4442 | -103.6528
9.4479 | -103.6734
9.4517 | -103.69
9.4554 | -103.7036
9.4591 | -103.7138
9.4628 | -103.7143
9.4666 | -103.7069
9.4703 | -103.6923
9.474 | -103.6782
9.4778 | -103.6602
9.4815 | -103.6383
9.4852 | -103.6096
9.4889 | -103.5786
9.4927 | -103.5456
9.4964 | -103.5096
9.5001 | -103.4751
9.5038 | -103.4428
9.5076 | -103.416
9.5113 | -103.3951
9.515 | -103.3759
9.5188 | -103.3638
9.5225 | -103.3553
9.5262 | -103.3492
9.5299 | -103.3496
9.5337 | -103.357
9.5374 | -103.3679
9.5411 | -103.3833
9.5448 | -103.4008
9.5486 | -103.4187
9.5523 | -103.4361
9.556 | -103.4575
9.5597 | -103.4701
9.5635 | -103.4826
9.5672 | -103.5016
9.5709 | -103.5156
9.5747 | -103.5268
9.5784 | -103.5325
9.5821 | -103.5352
9.5858 | -103.5358
9.5896 | -103.537
9.5933 | -103.5349
9.597 | -103.5341
9.6007 | -103.5335
9.6045 | -103.5306
9.6082 | -103.5225
9.6119 | -103.5115
9.6157 | -103.503
9.6194 | -103.4904
9.6231 | -103.4722
9.6268 | -103.4515
9.6306 | -103.4301
9.6343 | -103.4102
9.638 | -103.3841
9.6417 | -103.3546
9.6455 | -103.3196
9.6492 | -103.2817
9.6529 | -103.2418
9.6567 | -103.2064
9.6604 | -103.1758
9.6641 | -103.1451
9.6678 | -103.1152
9.6716 | -103.0879
9.6753 | -103.0829
9.679 | -103.0706
9.6827 | -103.0613
9.6865 | -103.0549
9.6902 | -103.0513
9.6939 | -103.0514
9.6976 | -103.0549
9.7014 | -103.0681
9.7051 | -103.0791
9.7088 | -103.0905
9.7126 | -103.1083
9.7163 | -103.1255
9.72 | -103.1509
9.7237 | -103.1802
9.7275 | -103.2155
9.7312 | -103.2534
9.7349 | -103.2967
9.7386 | -103.3404
9.7424 | -103.3855
9.7461 | -103.4324
9.7498 | -103.4848
9.7536 | -103.5348
9.7573 | -103.5847
9.761 | -103.6376
9.7647 | -103.6908
9.7685 | -103.7367
9.7722 | -103.7763
9.7759 | -103.8101
9.7796 | -103.8401
9.7834 | -103.8641
9.7871 | -103.8798
9.7908 | -103.8901
9.7945 | -103.8989
9.7983 | -103.8937
9.802 | -103.884
9.8057 | -103.8728
9.8095 | -103.856
9.8132 | -103.8352
9.8169 | -103.8072
9.8206 | -103.7771
9.8244 | -103.7472
9.8281 | -103.7185
9.8318 | -103.6932
9.8355 | -103.6725
9.8393 | -103.6553
9.843 | -103.6351
9.8467 | -103.6124
9.8505 | -103.5967
9.8542 | -103.5862
9.8579 | -103.5798
9.8616 | -103.5797
9.8654 | -103.5777
9.8691 | -103.5778
9.8728 | -103.5791
9.8765 | -103.5785
9.8803 | -103.5793
9.884 | -103.5831
9.8877 | -103.5817
9.8915 | -103.5796
9.8952 | -103.5762
9.8989 | -103.5707
9.9026 | -103.5637
9.9064 | -103.5566
9.9101 | -103.5536
9.9138 | -103.5531
9.9175 | -103.5528
9.9213 | -103.5506
9.925 | -103.5533
9.9287 | -103.5563
9.9324 | -103.5642
9.9362 | -103.572
9.9399 | -103.5787
9.9436 | -103.5894
9.9474 | -103.5997
9.9511 | -103.6164
9.9548 | -103.6354
9.9585 | -103.6514
9.9623 | -103.6694
9.966 | -103.6856
9.9697 | -103.6979
9.9734 | -103.7099
9.9772 | -103.7197
9.9809 | -103.7282
9.9846 | -103.7357
9.9884 | -103.746
9.9921 | -103.7526
9.9958 | -103.7548
9.9995 | -103.7545
10.0033 | -103.7478
10.007 | -103.7382
10.0107 | -103.7283
10.0144 | -103.7124
10.0182 | -103.7002
10.0219 | -103.6911
10.0256 | -103.6767
10.0294 | -103.6538
10.0331 | -103.6382
10.0368 | -103.6217
10.0405 | -103.6053
10.0443 | -103.5848
10.048 | -103.5652
10.0517 | -103.5499
10.0554 | -103.5347
10.0592 | -103.5221
10.0629 | -103.5157
10.0666 | -103.5086
10.0703 | -103.5046
10.0741 | -103.5052
10.0778 | -103.5101
10.0815 | -103.5215
10.0853 | -103.5375
10.089 | -103.5558
10.0927 | -103.5759
10.0964 | -103.5938
10.1002 | -103.6172
10.1039 | -103.6423
10.1076 | -103.6672
10.1113 | -103.6927
10.1151 | -103.7199
10.1188 | -103.7439
10.1225 | -103.758
10.1263 | -103.7669
10.13 | -103.7738
10.1337 | -103.7738
10.1374 | -103.771
10.1412 | -103.7651
10.1449 | -103.7564
10.1486 | -103.7477
10.1523 | -103.7355
10.1561 | -103.7206
10.1598 | -103.7065
10.1635 | -103.6837
10.1672 | -103.6605
10.171 | -103.6398
10.1747 | -103.6275
10.1784 | -103.6158
10.1822 | -103.6058
10.1859 | -103.5948
10.1896 | -103.593
10.1933 | -103.5936
10.1971 | -103.5974
10.2008 | -103.6068
10.2045 | -103.6165
10.2082 | -103.6289
10.212 | -103.6413
10.2157 | -103.6519
10.2194 | -103.6643
10.2232 | -103.6811
10.2269 | -103.6934
10.2306 | -103.703
10.2343 | -103.7122
10.2381 | -103.7207
10.2418 | -103.7325
10.2455 | -103.7441
10.2492 | -103.7531
10.253 | -103.7662
10.2567 | -103.7775
10.2604 | -103.7886
10.2642 | -103.8013
10.2679 | -103.8196
10.2716 | -103.8404
10.2753 | -103.8603
10.2791 | -103.8807
10.2828 | -103.9045
10.2865 | -103.9294
10.2902 | -103.9548
10.294 | -103.9748
10.2977 | -103.9954
10.3014 | -104.0159
10.3051 | -104.0329
10.3089 | -104.0475
10.3126 | -104.0618
10.3163 | -104.0767
10.3201 | -104.0889
10.3238 | -104.092
10.3275 | -104.0869
10.3312 | -104.0821
10.335 | -104.0823
10.3387 | -104.0785
10.3424 | -104.0756
10.3461 | -104.0688
10.3499 | -104.0622
10.3536 | -104.057
10.3573 | -104.0498
10.3611 | -104.0505
10.3648 | -104.0451
10.3685 | -104.0406
10.3722 | -104.0338
10.376 | -104.0226
10.3797 | -104.0111
10.3834 | -103.9986
10.3871 | -103.9864
10.3909 | -103.9788
10.3946 | -103.9635
10.3983 | -103.9486
10.402 | -103.9369
10.4058 | -103.9273
10.4095 | -103.9195
10.4132 | -103.9146
10.417 | -103.9082
10.4207 | -103.8993
10.4244 | -103.8972
10.4281 | -103.897
10.4319 | -103.8994
10.4356 | -103.8986
10.4393 | -103.8997
10.443 | -103.9014
10.4468 | -103.905
10.4505 | -103.9084
10.4542 | -103.9161
10.458 | -103.9284
10.4617 | -103.9405
10.4654 | -103.9518
10.4691 | -103.9616
10.4729 | -103.9764
10.4766 | -103.9948
10.4803 | -104.0152
10.484 | -104.0341
10.4878 | -104.0494
10.4915 | -104.0624
10.4952 | -104.0688
10.499 | -104.0773
10.5027 | -104.0876
10.5064 | -104.101
10.5101 | -104.1129
10.5139 | -104.1215
10.5176 | -104.1321
10.5213 | -104.1434
10.525 | -104.1527
10.5288 | -104.1657
10.5325 | -104.1822
10.5362 | -104.1973
10.5399 | -104.2075
10.5437 | -104.2156
10.5474 | -104.2289
10.5511 | -104.2406
10.5549 | -104.2493
10.5586 | -104.2542
10.5623 | -104.2579
10.566 | -104.2594
10.5698 | -104.2658
10.5735 | -104.2694
10.5772 | -104.2657
10.5809 | -104.256
10.5847 | -104.2454
10.5884 | -104.2347
10.5921 | -104.2183
10.5959 | -104.1998
10.5996 | -104.1818
10.6033 | -104.1595
10.607 | -104.1306
10.6108 | -104.0968
10.6145 | -104.064
10.6182 | -104.0321
10.6219 | -103.9988
10.6257 | -103.9641
10.6294 | -103.933
10.6331 | -103.9028
10.6369 | -103.8792
10.6406 | -103.8568
10.6443 | -103.8399
10.648 | -103.8279
10.6518 | -103.8121
10.6555 | -103.7951
10.6592 | -103.779
10.6629 | -103.7665
10.6667 | -103.7554
10.6704 | -103.7464
10.6741 | -103.7393
10.6778 | -103.7322
10.6816 | -103.7235
10.6853 | -103.7153
10.689 | -103.7111
10.6928 | -103.7081
10.6965 | -103.6997
10.7002 | -103.6906
10.7039 | -103.6833
10.7077 | -103.6758
10.7114 | -103.6699
10.7151 | -103.665
10.7188 | -103.6587
10.7226 | -103.6497
10.7263 | -103.6439
10.73 | -103.6413
10.7338 | -103.6418
10.7375 | -103.6408
10.7412 | -103.6449
10.7449 | -103.6508
10.7487 | -103.6577
10.7524 | -103.6715
10.7561 | -103.6795
10.7598 | -103.6872
10.7636 | -103.695
10.7673 | -103.703
10.771 | -103.713
10.7747 | -103.7238
10.7785 | -103.7355
10.7822 | -103.7444
10.7859 | -103.7514
10.7897 | -103.76
10.7934 | -103.7629
10.7971 | -103.768
10.8008 | -103.7715
10.8046 | -103.7753
10.8083 | -103.7798
10.812 | -103.787
10.8157 | -103.7966
10.8195 | -103.8067
10.8232 | -103.8154
10.8269 | -103.8246
10.8307 | -103.8354
10.8344 | -103.8467
10.8381 | -103.8581
10.8418 | -103.8665
10.8456 | -103.8728
10.8493 | -103.8795
10.853 | -103.8849
10.8567 | -103.8886
10.8605 | -103.8886
10.8642 | -103.8852
10.8679 | -103.8774
10.8717 | -103.8639
10.8754 | -103.8504
10.8791 | -103.8349
10.8828 | -103.8166
10.8866 | -103.7954
10.8903 | -103.7723
10.894 | -103.7476
10.8977 | -103.7241
10.9015 | -103.7046
10.9052 | -103.6862
10.9089 | -103.6701
10.9126 | -103.6587
10.9164 | -103.6488
10.9201 | -103.6437
10.9238 | -103.6409
10.9276 | -103.6386
10.9313 | -103.6339
10.935 | -103.6313
10.9387 | -103.6329
10.9425 | -103.6351
10.9462 | -103.6376
10.9499 | -103.6429
10.9536 | -103.6517
10.9574 | -103.6612
10.9611 | -103.6732
10.9648 | -103.6851
10.9686 | -103.7006
10.9723 | -103.7158
10.976 | -103.7303
10.9797 | -103.7465
10.9835 | -103.7681
10.9872 | -103.791
10.9909 | -103.8143
10.9946 | -103.8418
10.9984 | -103.8708
11.0021 | -103.9038
11.0058 | -103.9383
11.0096 | -103.9712
11.0133 | -104.0027
11.017 | -104.0339
11.0207 | -104.067
11.0245 | -104.0956
11.0282 | -104.1208
11.0319 | -104.1438
11.0356 | -104.1625
11.0394 | -104.1757
11.0431 | -104.1894
11.0468 | -104.1974
11.0505 | -104.1943
11.0543 | -104.1887
11.058 | -104.1812
11.0617 | -104.1641
11.0655 | -104.1431
11.0692 | -104.1161
11.0729 | -104.0862
11.0766 | -104.0545
11.0804 | -104.0183
11.0841 | -103.9801
11.0878 | -103.9449
11.0915 | -103.9112
11.0953 | -103.8766
11.099 | -103.8436
11.1027 | -103.8163
11.1065 | -103.7941
11.1102 | -103.7736
11.1139 | -103.7565
11.1176 | -103.7464
11.1214 | -103.74
11.1251 | -103.7411
11.1288 | -103.7467
11.1325 | -103.7565
11.1363 | -103.7711
11.14 | -103.7886
11.1437 | -103.8084
11.1474 | -103.8291
11.1512 | -103.8521
11.1549 | -103.8771
11.1586 | -103.9007
11.1624 | -103.9226
11.1661 | -103.9475
11.1698 | -103.9726
11.1735 | -103.9972
11.1773 | -104.0209
11.181 | -104.0402
11.1847 | -104.0616
11.1884 | -104.0775
11.1922 | -104.0928
11.1959 | -104.1059
11.1996 | -104.1166
11.2034 | -104.12
11.2071 | -104.1233
11.2108 | -104.1286
11.2145 | -104.1291
11.2183 | -104.1245
11.222 | -104.1228
11.2257 | -104.1214
11.2294 | -104.1193
11.2332 | -104.116
11.2369 | -104.1112
11.2406 | -104.1104
11.2444 | -104.1101
11.2481 | -104.1093
11.2518 | -104.1052
11.2555 | -104.1003
11.2593 | -104.0987
11.263 | -104.0987
11.2667 | -104.1013
11.2704 | -104.1049
11.2742 | -104.1126
11.2779 | -104.1168
11.2816 | -104.1221
11.2853 | -104.1299
11.2891 | -104.1388
11.2928 | -104.1445
11.2965 | -104.1496
11.3003 | -104.1543
11.304 | -104.1607
11.3077 | -104.17
11.3114 | -104.1785
11.3152 | -104.1863
11.3189 | -104.1918
11.3226 | -104.1974
11.3263 | -104.2049
11.3301 | -104.2116
11.3338 | -104.2219
11.3375 | -104.2349
11.3413 | -104.2448
11.345 | -104.2561
11.3487 | -104.2685
11.3524 | -104.2857
11.3562 | -104.3003
11.3599 | -104.3152
11.3636 | -104.3306
11.3673 | -104.3503
11.3711 | -104.371
11.3748 | -104.3899
11.3785 | -104.4084
11.3823 | -104.4274
11.386 | -104.4458
11.3897 | -104.4622
11.3934 | -104.48
11.3972 | -104.4974
11.4009 | -104.5146
11.4046 | -104.5261
11.4083 | -104.5366
11.4121 | -104.5463
11.4158 | -104.5597
11.4195 | -104.5702
11.4232 | -104.5765
11.427 | -104.5811
11.4307 | -104.5814
11.4344 | -104.5786
11.4382 | -104.5713
11.4419 | -104.5639
11.4456 | -104.5582
11.4493 | -104.5492
11.4531 | -104.5384
11.4568 | -104.5296
11.4605 | -104.5208
11.4642 | -104.5097
11.468 | -104.4971
11.4717 | -104.4851
11.4754 | -104.4729
11.4792 | -104.467
11.4829 | -104.4645
11.4866 | -104.4648
11.4903 | -104.4673
11.4941 | -104.4702
11.4978 | -104.4713
11.5015 | -104.4748
11.5052 | -104.4794
11.509 | -104.4887
11.5127 | -104.4993
11.5164 | -104.5099
11.5201 | -104.5208
11.5239 | -104.5325
11.5276 | -104.5406
11.5313 | -104.5484
11.5351 | -104.5512
11.5388 | -104.5503
11.5425 | -104.5498
11.5462 | -104.5479
11.55 | -104.5444
11.5537 | -104.5408
11.5574 | -104.536
11.5611 | -104.5295
11.5649 | -104.5219
11.5686 | -104.5104
11.5723 | -104.5014
11.5761 | -104.4924
11.5798 | -104.4821
11.5835 | -104.4718
11.5872 | -104.4613
11.591 | -104.4483
11.5947 | -104.4368
11.5984 | -104.4277
11.6021 | -104.4155
11.6059 | -104.4062
11.6096 | -104.3998
11.6133 | -104.3909
11.6171 | -104.384
11.6208 | -104.3773
11.6245 | -104.3726
11.6282 | -104.368
11.632 | -104.3625
11.6357 | -104.357
11.6394 | -104.3485
11.6431 | -104.3376
11.6469 | -104.3235
11.6506 | -104.306
11.6543 | -104.287
11.658 | -104.2728
11.6618 | -104.2585
11.6655 | -104.2396
11.6692 | -104.2161
11.673 | -104.1956
11.6767 | -104.1752
11.6804 | -104.155
11.6841 | -104.1359
11.6879 | -104.1165
11.6916 | -104.1006
11.6953 | -104.0882
11.699 | -104.078
11.7028 | -104.0722
11.7065 | -104.0713
11.7102 | -104.071
11.714 | -104.0711
11.7177 | -104.0757
11.7214 | -104.0845
11.7251 | -104.096
11.7289 | -104.1066
11.7326 | -104.1204
11.7363 | -104.1368
11.74 | -104.1548
11.7438 | -104.1745
11.7475 | -104.1958
11.7512 | -104.2184
11.7549 | -104.2458
11.7587 | -104.2777
11.7624 | -104.3105
11.7661 | -104.3405
11.7699 | -104.372
11.7736 | -104.4033
11.7773 | -104.4352
11.781 | -104.4661
11.7848 | -104.4915
11.7885 | -104.5194
11.7922 | -104.5442
11.7959 | -104.5663
11.7997 | -104.5911
11.8034 | -104.614
11.8071 | -104.6336
11.8109 | -104.6518
11.8146 | -104.6678
11.8183 | -104.6794
11.822 | -104.6914
11.8258 | -104.6963
11.8295 | -104.7032
11.8332 | -104.7082
11.8369 | -104.7173
11.8407 | -104.7277
11.8444 | -104.7331
11.8481 | -104.7375
11.8519 | -104.7414
11.8556 | -104.7432
11.8593 | -104.7454
11.863 | -104.7428
11.8668 | -104.7371
11.8705 | -104.7364
11.8742 | -104.7366
11.8779 | -104.734
11.8817 | -104.7301
11.8854 | -104.7273
11.8891 | -104.7217
11.8928 | -104.7132
11.8966 | -104.7088
11.9003 | -104.7093
11.904 | -104.7053
11.9078 | -104.7002
11.9115 | -104.6962
11.9152 | -104.687
11.9189 | -104.6772
11.9227 | -104.6657
11.9264 | -104.6561
11.9301 | -104.6445
11.9338 | -104.634
11.9376 | -104.6253
11.9413 | -104.6154
11.945 | -104.602
11.9488 | -104.588
11.9525 | -104.577
11.9562 | -104.5664
11.9599 | -104.5542
11.9637 | -104.543
11.9674 | -104.5334
11.9711 | -104.5222
11.9748 | -104.5104
11.9786 | -104.5011
11.9823 | -104.4921
11.986 | -104.4854
11.9898 | -104.479
11.9935 | -104.4693
11.9972 | -104.4592
12.0009 | -104.4534
12.0047 | -104.449
12.0084 | -104.4472
12.0121 | -104.4453
12.0158 | -104.4415
12.0196 | -104.4375
12.0233 | -104.434
12.027 | -104.435
12.0307 | -104.4429
12.0345 | -104.4548
12.0382 | -104.4687
12.0419 | -104.4818
12.0457 | -104.4893
12.0494 | -104.497
12.0531 | -104.5064
12.0568 | -104.5176
12.0606 | -104.5288
12.0643 | -104.5382
12.068 | -104.5501
12.0717 | -104.5567
12.0755 | -104.5597
12.0792 | -104.5651
12.0829 | -104.568
12.0867 | -104.5662
12.0904 | -104.565
12.0941 | -104.5638
12.0978 | -104.5613
12.1016 | -104.5532
12.1053 | -104.5446
12.109 | -104.5334
12.1127 | -104.5197
12.1165 | -104.5089
12.1202 | -104.4958
12.1239 | -104.4793
12.1276 | -104.4636
12.1314 | -104.4486
12.1351 | -104.4344
12.1388 | -104.4202
12.1426 | -104.4054
12.1463 | -104.3919
12.15 | -104.3799
12.1537 | -104.3683
12.1575 | -104.3571
12.1612 | -104.3478
12.1649 | -104.3428
12.1686 | -104.3385
12.1724 | -104.3362
12.1761 | -104.3345
12.1798 | -104.3322
12.1836 | -104.3316
12.1873 | -104.3388
12.191 | -104.3512
12.1947 | -104.367
12.1985 | -104.3847
12.2022 | -104.4025
12.2059 | -104.4185
12.2096 | -104.4343
12.2134 | -104.4518
12.2171 | -104.4699
12.2208 | -104.4907
12.2246 | -104.511
12.2283 | -104.5306
12.232 | -104.5484
12.2357 | -104.5608
12.2395 | -104.5735
12.2432 | -104.5833
12.2469 | -104.5915
12.2506 | -104.6
12.2544 | -104.6043
12.2581 | -104.6071
12.2618 | -104.6073
12.2655 | -104.6034
12.2693 | -104.5932
12.273 | -104.5836
12.2767 | -104.5729
12.2805 | -104.5615
12.2842 | -104.5503
12.2879 | -104.54
12.2916 | -104.5241
12.2954 | -104.5075
12.2991 | -104.4948
12.3028 | -104.481
12.3065 | -104.471
12.3103 | -104.4592
12.314 | -104.4509
12.3177 | -104.4449
12.3215 | -104.4417
12.3252 | -104.442
12.3289 | -104.445
12.3326 | -104.4499
12.3364 | -104.4544
12.3401 | -104.4632
12.3438 | -104.4746
12.3475 | -104.4854
12.3513 | -104.4968
12.355 | -104.5097
12.3587 | -104.5271
12.3625 | -104.5406
12.3662 | -104.5544
12.3699 | -104.567
12.3736 | -104.5766
12.3774 | -104.5841
12.3811 | -104.5946
12.3848 | -104.6073
12.3885 | -104.6186
12.3923 | -104.6297
12.396 | -104.642
12.3997 | -104.6464
12.4034 | -104.6475
12.4072 | -104.6468
12.4109 | -104.6453
12.4146 | -104.6443
12.4184 | -104.6386
12.4221 | -104.6288
12.4258 | -104.6176
12.4295 | -104.6012
12.4333 | -104.5843
12.437 | -104.5611
12.4407 | -104.5401
12.4444 | -104.5198
12.4482 | -104.4995
12.4519 | -104.48
12.4556 | -104.4609
12.4594 | -104.4432
12.4631 | -104.4245
12.4668 | -104.4071
12.4705 | -104.3935
12.4743 | -104.3807
12.478 | -104.3672
12.4817 | -104.3581
12.4854 | -104.3503
12.4892 | -104.3456
12.4929 | -104.3435
12.4966 | -104.3394
12.5003 | -104.3362
12.5041 | -104.3335
12.5078 | -104.3336
12.5115 | -104.3386
12.5153 | -104.3438
12.519 | -104.3459
12.5227 | -104.3503
12.5264 | -104.3582
12.5302 | -104.3658
12.5339 | -104.373
12.5376 | -104.3764
12.5413 | -104.3795
12.5451 | -104.3869
12.5488 | -104.3918
12.5525 | -104.3957
12.5563 | -104.4014
12.56 | -104.4044
12.5637 | -104.4054
12.5674 | -104.4066
12.5712 | -104.4086
12.5749 | -104.4119
12.5786 | -104.4147
12.5823 | -104.4193
12.5861 | -104.4208
12.5898 | -104.421
12.5935 | -104.4188
12.5973 | -104.4211
12.601 | -104.4247
12.6047 | -104.4269
12.6084 | -104.4317
12.6122 | -104.435
12.6159 | -104.4355
12.6196 | -104.4406
12.6233 | -104.4416
12.6271 | -104.4404
12.6308 | -104.4421
12.6345 | -104.4413
12.6382 | -104.4409
12.642 | -104.4389
12.6457 | -104.4375
12.6494 | -104.4345
12.6532 | -104.4347
12.6569 | -104.4365
12.6606 | -104.4379
12.6643 | -104.4392
12.6681 | -104.4418
12.6718 | -104.4439
12.6755 | -104.4434
12.6792 | -104.448
12.683 | -104.4565
12.6867 | -104.4656
12.6904 | -104.472
12.6942 | -104.4771
12.6979 | -104.4843
12.7016 | -104.492
12.7053 | -104.4986
12.7091 | -104.5055
12.7128 | -104.5137
12.7165 | -104.5241
12.7202 | -104.5353
12.724 | -104.5458
12.7277 | -104.5546
12.7314 | -104.5667
12.7352 | -104.5784
12.7389 | -104.5889
12.7426 | -104.6026
12.7463 | -104.6118
12.7501 | -104.6197
12.7538 | -104.6283
12.7575 | -104.6327
12.7612 | -104.6357
12.765 | -104.6417
12.7687 | -104.6462
12.7724 | -104.6501
12.7761 | -104.6546
12.7799 | -104.6533
12.7836 | -104.6528
12.7873 | -104.6544
12.7911 | -104.6545
12.7948 | -104.6558
12.7985 | -104.6583
12.8022 | -104.6587
12.806 | -104.6579
12.8097 | -104.6628
12.8134 | -104.6651
12.8171 | -104.6671
12.8209 | -104.6712
12.8246 | -104.6779
12.8283 | -104.6823
12.8321 | -104.6884
12.8358 | -104.6936
12.8395 | -104.6961
12.8432 | -104.7014
12.847 | -104.706
12.8507 | -104.7091
12.8544 | -104.7161
12.8581 | -104.7271
12.8619 | -104.7359
12.8656 | -104.7405
12.8693 | -104.7489
12.873 | -104.7584
12.8768 | -104.7703
12.8805 | -104.7842
12.8842 | -104.7983
12.888 | -104.8122
12.8917 | -104.8239
12.8954 | -104.8358
12.8991 | -104.8471
12.9029 | -104.8572
12.9066 | -104.8684
12.9103 | -104.882
12.914 | -104.896
12.9178 | -104.9033
12.9215 | -104.9102
12.9252 | -104.9193
12.929 | -104.9276
12.9327 | -104.9371
12.9364 | -104.9449
12.9401 | -104.9517
12.9439 | -104.9597
12.9476 | -104.9669
12.9513 | -104.9693
12.955 | -104.9743
12.9588 | -104.9786
12.9625 | -104.9813
12.9662 | -104.9823
12.97 | -104.983
12.9737 | -104.9837
12.9774 | -104.9863
12.9811 | -104.9902
12.9849 | -104.9903
12.9886 | -104.99
12.9923 | -104.9902
12.996 | -104.9885
12.9998 | -104.9846
13.0035 | -104.9799
13.0072 | -104.9756
13.0109 | -104.9693
13.0147 | -104.9626
13.0184 | -104.9525
13.0221 | -104.9397
13.0259 | -104.9281
13.0296 | -104.9141
13.0333 | -104.8982
13.037 | -104.8822
13.0408 | -104.8609
13.0445 | -104.8367
13.0482 | -104.8187
13.0519 | -104.7997
13.0557 | -104.7772
13.0594 | -104.7566
13.0631 | -104.7345
13.0669 | -104.7111
13.0706 | -104.6904
13.0743 | -104.67
13.078 | -104.6471
13.0818 | -104.6288
13.0855 | -104.6112
13.0892 | -104.5953
13.0929 | -104.5805
13.0967 | -104.5692
13.1004 | -104.5612
13.1041 | -104.5533
13.1078 | -104.5505
13.1116 | -104.5494
13.1153 | -104.5487
13.119 | -104.5475
13.1228 | -104.5499
13.1265 | -104.558
13.1302 | -104.5705
13.1339 | -104.5852
13.1377 | -104.6029
13.1414 | -104.6167
13.1451 | -104.6305
13.1488 | -104.6452
13.1526 | -104.6605
13.1563 | -104.6773
13.16 | -104.694
13.1638 | -104.7092
13.1675 | -104.7222
13.1712 | -104.7327
13.1749 | -104.7449
13.1787 | -104.7509
13.1824 | -104.754
13.1861 | -104.7519
13.1898 | -104.7516
13.1936 | -104.7507
13.1973 | -104.7428
13.201 | -104.7313
13.2048 | -104.722
13.2085 | -104.7075
13.2122 | -104.6907
13.2159 | -104.6722
13.2197 | -104.6503
13.2234 | -104.6326
13.2271 | -104.6142
13.2308 | -104.5966
13.2346 | -104.5783
13.2383 | -104.5579
13.242 | -104.542
13.2457 | -104.5276
13.2495 | -104.5119
13.2532 | -104.5001
13.2569 | -104.4913
13.2607 | -104.4804
13.2644 | -104.471
13.2681 | -104.4636
13.2718 | -104.4541
13.2756 | -104.4498
13.2793 | -104.4463
13.283 | -104.4436
13.2867 | -104.4414
13.2905 | -104.4392
13.2942 | -104.4386
13.2979 | -104.4373
13.3017 | -104.4379
13.3054 | -104.4425
13.3091 | -104.4455
13.3128 | -104.4431
13.3166 | -104.4373
13.3203 | -104.4279
13.324 | -104.4191
13.3277 | -104.41
13.3315 | -104.4041
13.3352 | -104.3961
13.3389 | -104.383
13.3427 | -104.3674
13.3464 | -104.3519
13.3501 | -104.3331
13.3538 | -104.3111
13.3576 | -104.2868
13.3613 | -104.263
13.365 | -104.2393
13.3687 | -104.2167
13.3725 | -104.1896
13.3762 | -104.1653
13.3799 | -104.1374
13.3836 | -104.1137
13.3874 | -104.0948
13.3911 | -104.0776
13.3948 | -104.0617
13.3986 | -104.0512
13.4023 | -104.0412
13.406 | -104.036
13.4097 | -104.0362
13.4135 | -104.0379
13.4172 | -104.043
13.4209 | -104.0505
13.4246 | -104.0551
13.4284 | -104.0657
13.4321 | -104.0821
13.4358 | -104.1007
13.4396 | -104.1235
13.4433 | -104.1418
13.447 | -104.1571
13.4507 | -104.1725
13.4545 | -104.1881
13.4582 | -104.2026
13.4619 | -104.2193
13.4656 | -104.2343
13.4694 | -104.2473
13.4731 | -104.2601
13.4768 | -104.2733
13.4805 | -104.2842
13.4843 | -104.2924
13.488 | -104.303
13.4917 | -104.3158
13.4955 | -104.326
13.4992 | -104.3325
13.5029 | -104.3411
13.5066 | -104.3513
13.5104 | -104.363
13.5141 | -104.3764
13.5178 | -104.3885
13.5215 | -104.3991
13.5253 | -104.4089
13.529 | -104.4186
13.5327 | -104.4306
13.5365 | -104.4409
13.5402 | -104.4508
13.5439 | -104.4607
13.5476 | -104.4713
13.5514 | -104.4805
13.5551 | -104.4891
13.5588 | -104.4959
13.5625 | -104.5036
13.5663 | -104.5159
13.57 | -104.5259
13.5737 | -104.5369
13.5775 | -104.5414
13.5812 | -104.5474
13.5849 | -104.5558
13.5886 | -104.5596
13.5924 | -104.5636
13.5961 | -104.5679
13.5998 | -104.5695
13.6035 | -104.5723
13.6073 | -104.5721
13.611 | -104.5721
13.6147 | -104.5755
13.6184 | -104.5811
13.6222 | -104.5778
13.6259 | -104.5717
13.6296 | -104.567
13.6334 | -104.5644
13.6371 | -104.562
13.6408 | -104.5594
13.6445 | -104.5571
13.6483 | -104.556
13.652 | -104.5549
13.6557 | -104.5511
13.6594 | -104.5454
13.6632 | -104.5413
13.6669 | -104.5362
13.6706 | -104.5328
13.6744 | -104.5295
13.6781 | -104.5292
13.6818 | -104.532
13.6855 | -104.5326
13.6893 | -104.5325
13.693 | -104.5332
13.6967 | -104.5345
13.7004 | -104.537
13.7042 | -104.5387
13.7079 | -104.5343
13.7116 | -104.5327
13.7154 | -104.5313
13.7191 | -104.5299
13.7228 | -104.5283
13.7265 | -104.5242
13.7303 | -104.5213
13.734 | -104.5227
13.7377 | -104.5221
13.7414 | -104.5197
13.7452 | -104.5185
13.7489 | -104.5207
13.7526 | -104.5237
13.7563 | -104.5256
13.7601 | -104.5253
13.7638 | -104.5233
13.7675 | -104.5204
13.7713 | -104.5193
13.775 | -104.5206
13.7787 | -104.5196
13.7824 | -104.5172
13.7862 | -104.5171
13.7899 | -104.5155
13.7936 | -104.5142
13.7973 | -104.5137
13.8011 | -104.5099
13.8048 | -104.5092
13.8085 | -104.5112
13.8123 | -104.513
13.816 | -104.5176
13.8197 | -104.5241
13.8234 | -104.5305
13.8272 | -104.5389
13.8309 | -104.5509
13.8346 | -104.5649
13.8383 | -104.5798
13.8421 | -104.5987
13.8458 | -104.6158
13.8495 | -104.6346
13.8532 | -104.6568
13.857 | -104.679
13.8607 | -104.6997
13.8644 | -104.7202
13.8682 | -104.7395
13.8719 | -104.7584
13.8756 | -104.7763
13.8793 | -104.794
13.8831 | -104.8091
13.8868 | -104.8218
13.8905 | -104.8356
13.8942 | -104.8452
13.898 | -104.8522
13.9017 | -104.8579
13.9054 | -104.8664
13.9092 | -104.8671
13.9129 | -104.8637
13.9166 | -104.8595
13.9203 | -104.8551
13.9241 | -104.8492
13.9278 | -104.8414
13.9315 | -104.83
13.9352 | -104.8198
13.939 | -104.8116
13.9427 | -104.8003
13.9464 | -104.7852
13.9502 | -104.7702
13.9539 | -104.7568
13.9576 | -104.7442
13.9613 | -104.7238
13.9651 | -104.6974
13.9688 | -104.6747
13.9725 | -104.6549
13.9762 | -104.6342
13.98 | -104.6148
13.9837 | -104.5976
13.9874 | -104.5814
13.9911 | -104.5621
13.9949 | -104.5409
13.9986 | -104.5209
14.0023 | -104.5032
14.0061 | -104.4849
14.0098 | -104.464
14.0135 | -104.443
14.0172 | -104.4239
14.021 | -104.4034
14.0247 | -104.3834
14.0284 | -104.3674
14.0321 | -104.3523
14.0359 | -104.3372
14.0396 | -104.3246
14.0433 | -104.3145
14.0471 | -104.3081
14.0508 | -104.3021
14.0545 | -104.2991
14.0582 | -104.3008
14.062 | -104.3061
14.0657 | -104.314
14.0694 | -104.3235
14.0731 | -104.3352
14.0769 | -104.3487
14.0806 | -104.3611
14.0843 | -104.3773
14.0881 | -104.3941
14.0918 | -104.411
14.0955 | -104.4258
14.0992 | -104.4377
14.103 | -104.4493
14.1067 | -104.4607
14.1104 | -104.4711
14.1141 | -104.4791
14.1179 | -104.4843
14.1216 | -104.4878
14.1253 | -104.4891
14.129 | -104.4838
14.1328 | -104.4771
14.1365 | -104.469
14.1402 | -104.4603
14.144 | -104.4516
14.1477 | -104.4419
14.1514 | -104.4318
14.1551 | -104.4191
14.1589 | -104.4068
14.1626 | -104.4006
14.1663 | -104.398
14.17 | -104.3883
14.1738 | -104.3769
14.1775 | -104.3661
14.1812 | -104.3597
14.185 | -104.3572
14.1887 | -104.3556
14.1924 | -104.3582
14.1961 | -104.3676
14.1999 | -104.3799
14.2036 | -104.3928
14.2073 | -104.4085
14.211 | -104.4247
14.2148 | -104.4465
14.2185 | -104.4717
14.2222 | -104.498
14.2259 | -104.5299
14.2297 | -104.5639
14.2334 | -104.5965
14.2371 | -104.6272
14.2409 | -104.6583
14.2446 | -104.6931
14.2483 | -104.7251
14.252 | -104.7565
14.2558 | -104.7895
14.2595 | -104.8178
14.2632 | -104.8434
14.2669 | -104.8656
14.2707 | -104.8876
14.2744 | -104.9099
14.2781 | -104.9254
14.2819 | -104.9363
14.2856 | -104.9499
14.2893 | -104.9577
14.293 | -104.959
14.2968 | -104.955
14.3005 | -104.9543
14.3042 | -104.9512
14.3079 | -104.9419
14.3117 | -104.9271
14.3154 | -104.9117
14.3191 | -104.894
14.3229 | -104.8744
14.3266 | -104.855
14.3303 | -104.8365
14.334 | -104.8191
14.3378 | -104.7973
14.3415 | -104.7735
14.3452 | -104.7533
14.3489 | -104.7367
14.3527 | -104.7216
14.3564 | -104.7062
14.3601 | -104.6923
14.3638 | -104.6821
14.3676 | -104.6731
14.3713 | -104.6646
14.375 | -104.6607
14.3788 | -104.6582
14.3825 | -104.6562
14.3862 | -104.652
14.3899 | -104.6473
14.3937 | -104.6458
14.3974 | -104.642
14.4011 | -104.6409
14.4048 | -104.6427
14.4086 | -104.6441
14.4123 | -104.6454
14.416 | -104.6479
14.4198 | -104.6494
14.4235 | -104.6504
14.4272 | -104.6511
14.4309 | -104.65
14.4347 | -104.6486
14.4384 | -104.6447
14.4421 | -104.6434
14.4458 | -104.6452
14.4496 | -104.6422
14.4533 | -104.6396
14.457 | -104.6357
14.4608 | -104.6317
14.4645 | -104.6289
14.4682 | -104.6258
14.4719 | -104.6219
14.4757 | -104.6184
14.4794 | -104.6164
14.4831 | -104.6206
14.4868 | -104.6256
14.4906 | -104.6255
14.4943 | -104.6286
14.498 | -104.6365
14.5017 | -104.645
14.5055 | -104.6524
14.5092 | -104.6616
14.5129 | -104.6733
14.5167 | -104.6881
14.5204 | -104.701
14.5241 | -104.7146
14.5278 | -104.7286
14.5316 | -104.7428
14.5353 | -104.754
14.539 | -104.7638
14.5427 | -104.7781
14.5465 | -104.7936
14.5502 | -104.8099
14.5539 | -104.8208
14.5577 | -104.8266
14.5614 | -104.8328
14.5651 | -104.8401
14.5688 | -104.8468
14.5726 | -104.8512
14.5763 | -104.855
14.58 | -104.8577
14.5837 | -104.8565
14.5875 | -104.8544
14.5912 | -104.852
14.5949 | -104.8493
14.5986 | -104.8464
14.6024 | -104.8434
14.6061 | -104.8409
14.6098 | -104.8327
14.6136 | -104.8245
14.6173 | -104.8166
14.621 | -104.8057
14.6247 | -104.7946
14.6285 | -104.7835
14.6322 | -104.7717
14.6359 | -104.7582
14.6396 | -104.7455
14.6434 | -104.7335
14.6471 | -104.7217
14.6508 | -104.707
14.6546 | -104.6943
14.6583 | -104.6799
14.662 | -104.6653
14.6657 | -104.6551
14.6695 | -104.6429
14.6732 | -104.6337
14.6769 | -104.6251
14.6806 | -104.619
14.6844 | -104.6098
14.6881 | -104.6004
14.6918 | -104.5915
14.6956 | -104.5864
14.6993 | -104.5833
14.703 | -104.5807
14.7067 | -104.5794
14.7105 | -104.58
14.7142 | -104.5826
14.7179 | -104.5851
14.7216 | -104.5864
14.7254 | -104.5883
14.7291 | -104.5904
14.7328 | -104.5931
14.7365 | -104.5969
14.7403 | -104.6034
14.744 | -104.6121
14.7477 | -104.6215
14.7515 | -104.6317
14.7552 | -104.6391
14.7589 | -104.6461
14.7626 | -104.6482
14.7664 | -104.6558
14.7701 | -104.6656
14.7738 | -104.6726
14.7775 | -104.6811
14.7813 | -104.6928
14.785 | -104.7019
14.7887 | -104.7095
14.7925 | -104.7167
14.7962 | -104.7238
14.7999 | -104.7281
14.8036 | -104.7289
14.8074 | -104.7293
14.8111 | -104.7333
14.8148 | -104.7341
14.8185 | -104.7331
14.8223 | -104.7347
14.826 | -104.7353
14.8297 | -104.7352
14.8334 | -104.7335
14.8372 | -104.7291
14.8409 | -104.7244
14.8446 | -104.7194
14.8484 | -104.7168
14.8521 | -104.7136
14.8558 | -104.7082
14.8595 | -104.7032
14.8633 | -104.6993
14.867 | -104.6924
14.8707 | -104.6881
14.8744 | -104.6827
14.8782 | -104.6767
14.8819 | -104.6733
14.8856 | -104.6676
14.8894 | -104.6623
14.8931 | -104.659
14.8968 | -104.6574
14.9005 | -104.6569
14.9043 | -104.6536
14.908 | -104.6504
14.9117 | -104.6496
14.9154 | -104.6488
14.9192 | -104.645
14.9229 | -104.6427
14.9266 | -104.6413
14.9304 | -104.6398
14.9341 | -104.639
14.9378 | -104.6405
14.9415 | -104.6429
14.9453 | -104.6411
14.949 | -104.6396
14.9527 | -104.6373
14.9564 | -104.6333
14.9602 | -104.6331
14.9639 | -104.6328
14.9676 | -104.6325
14.9713 | -104.6353
14.9751 | -104.6432
14.9788 | -104.6442
14.9825 | -104.6455
14.9863 | -104.6495
14.99 | -104.6518
14.9937 | -104.6529
14.9974 | -104.6554
15.0012 | -104.6601
15.0049 | -104.6681
15.0086 | -104.6772
15.0123 | -104.6856
15.0161 | -104.6978
15.0198 | -104.7098
15.0235 | -104.7154
15.0273 | -104.7254
15.031 | -104.7382
15.0347 | -104.7504
15.0384 | -104.761
15.0422 | -104.7711
15.0459 | -104.785
15.0496 | -104.8021
15.0533 | -104.812
15.0571 | -104.8226
15.0608 | -104.8315
15.0645 | -104.839
15.0683 | -104.8455
15.072 | -104.8518
15.0757 | -104.8576
15.0794 | -104.8658
15.0832 | -104.8722
15.0869 | -104.8772
15.0906 | -104.8804
15.0943 | -104.8799
15.0981 | -104.8773
15.1018 | -104.8765
15.1055 | -104.8763
15.1092 | -104.8765
15.113 | -104.8764
15.1167 | -104.8759
15.1204 | -104.8752
15.1242 | -104.8703
15.1279 | -104.865
15.1316 | -104.8574
15.1353 | -104.8534
15.1391 | -104.8498
15.1428 | -104.8445
15.1465 | -104.8414
15.1502 | -104.8386
15.154 | -104.8343
15.1577 | -104.8283
15.1614 | -104.825
15.1652 | -104.8208
15.1689 | -104.8178
15.1726 | -104.8158
15.1763 | -104.8117
15.1801 | -104.8084
15.1838 | -104.8077
15.1875 | -104.8065
15.1912 | -104.8026
15.195 | -104.801
15.1987 | -104.7961
15.2024 | -104.7933
15.2061 | -104.7916
15.2099 | -104.7895
15.2136 | -104.7856
15.2173 | -104.7802
15.2211 | -104.7738
15.2248 | -104.7682
15.2285 | -104.765
15.2322 | -104.7584
15.236 | -104.7519
15.2397 | -104.7459
15.2434 | -104.7368
15.2471 | -104.7276
15.2509 | -104.717
15.2546 | -104.7087
15.2583 | -104.7029
15.2621 | -104.6951
15.2658 | -104.6839
15.2695 | -104.6725
15.2732 | -104.6632
15.277 | -104.6546
15.2807 | -104.6481
15.2844 | -104.6416
15.2881 | -104.6359
15.2919 | -104.6309
15.2956 | -104.6242
15.2993 | -104.6178
15.3031 | -104.6117
15.3068 | -104.6091
15.3105 | -104.6098
15.3142 | -104.612
15.318 | -104.6156
15.3217 | -104.6188
15.3254 | -104.6193
15.3291 | -104.6205
15.3329 | -104.6254
15.3366 | -104.6312
15.3403 | -104.6405
15.344 | -104.65
15.3478 | -104.659
15.3515 | -104.6681
15.3552 | -104.6743
15.359 | -104.6778
15.3627 | -104.6825
15.3664 | -104.6865
15.3701 | -104.6906
15.3739 | -104.6919
15.3776 | -104.6887
15.3813 | -104.685
15.385 | -104.6823
15.3888 | -104.6782
15.3925 | -104.6719
15.3962 | -104.6644
15.4 | -104.6564
15.4037 | -104.6456
15.4074 | -104.6354
15.4111 | -104.6259
15.4149 | -104.6147
15.4186 | -104.6055
15.4223 | -104.5905
15.426 | -104.5739
15.4298 | -104.5602
15.4335 | -104.5496
15.4372 | -104.5388
15.441 | -104.5272
15.4447 | -104.5157
15.4484 | -104.5032
15.4521 | -104.4929
15.4559 | -104.4816
15.4596 | -104.4712
15.4633 | -104.4626
15.467 | -104.456
15.4708 | -104.4514
15.4745 | -104.4465
15.4782 | -104.4445
15.4819 | -104.4456
15.4857 | -104.4509
15.4894 | -104.4575
15.4931 | -104.4632
15.4969 | -104.4672
15.5006 | -104.4717
15.5043 | -104.4774
15.508 | -104.485
15.5118 | -104.4946
15.5155 | -104.5057
15.5192 | -104.5153
15.5229 | -104.5278
15.5267 | -104.5385
15.5304 | -104.5502
15.5341 | -104.5653
15.5379 | -104.5781
15.5416 | -104.5906
15.5453 | -104.6063
15.549 | -104.6217
15.5528 | -104.6352
15.5565 | -104.6508
15.5602 | -104.6656
15.5639 | -104.6832
15.5677 | -104.6972
15.5714 | -104.7136
15.5751 | -104.7275
15.5788 | -104.7427
15.5826 | -104.756
15.5863 | -104.7673
15.59 | -104.7757
15.5938 | -104.7845
15.5975 | -104.7919
15.6012 | -104.8029
15.6049 | -104.8092
15.6087 | -104.8158
15.6124 | -104.821
15.6161 | -104.8231
15.6198 | -104.8238
15.6236 | -104.823
15.6273 | -104.8223
15.631 | -104.82
15.6348 | -104.8159
15.6385 | -104.8087
15.6422 | -104.8013
15.6459 | -104.7956
15.6497 | -104.7901
15.6534 | -104.7801
15.6571 | -104.7679
15.6608 | -104.7568
15.6646 | -104.7459
15.6683 | -104.7323
15.672 | -104.7214
15.6758 | -104.7094
15.6795 | -104.6985
15.6832 | -104.6871
15.6869 | -104.6763
15.6907 | -104.6661
15.6944 | -104.6572
15.6981 | -104.6487
15.7018 | -104.6393
15.7056 | -104.6318
15.7093 | -104.6242
15.713 | -104.6205
15.7167 | -104.6201
15.7205 | -104.6187
15.7242 | -104.6187
15.7279 | -104.6217
15.7317 | -104.6233
15.7354 | -104.626
15.7391 | -104.63
15.7428 | -104.6339
15.7466 | -104.6391
15.7503 | -104.6459
15.754 | -104.6485
15.7577 | -104.6495
15.7615 | -104.654
15.7652 | -104.6575
15.7689 | -104.6612
15.7727 | -104.6636
15.7764 | -104.6654
15.7801 | -104.6628
15.7838 | -104.6605
15.7876 | -104.6564
15.7913 | -104.6522
15.795 | -104.647
15.7987 | -104.642
15.8025 | -104.6326
15.8062 | -104.622
15.8099 | -104.612
15.8137 | -104.601
15.8174 | -104.5884
15.8211 | -104.5787
15.8248 | -104.5688
15.8286 | -104.5563
15.8323 | -104.541
15.836 | -104.5241
15.8397 | -104.5091
15.8435 | -104.4961
15.8472 | -104.4826
15.8509 | -104.47
15.8546 | -104.4613
15.8584 | -104.4519
15.8621 | -104.443
15.8658 | -104.4392
15.8696 | -104.4363
15.8733 | -104.4333
15.877 | -104.4309
15.8807 | -104.4318
15.8845 | -104.436
15.8882 | -104.4426
15.8919 | -104.4518
15.8956 | -104.4637
15.8994 | -104.4742
15.9031 | -104.4877
15.9068 | -104.5012
15.9106 | -104.5152
15.9143 | -104.5274
15.918 | -104.5398
15.9217 | -104.5514
15.9255 | -104.5636
15.9292 | -104.5769
15.9329 | -104.5885
15.9366 | -104.5993
15.9404 | -104.6088
15.9441 | -104.6166
15.9478 | -104.6283
15.9515 | -104.6356
15.9553 | -104.6355
15.959 | -104.6343
15.9627 | -104.6329
15.9665 | -104.6336
15.9702 | -104.6325
15.9739 | -104.6279
15.9776 | -104.6219
15.9814 | -104.6154
15.9851 | -104.6075
15.9888 | -104.5979
15.9925 | -104.5869
15.9963 | -104.5737
16 | -104.5605
16.0037 | -104.5461
16.0075 | -104.5337
16.0112 | -104.5215
16.0149 | -104.5096
16.0186 | -104.4986
16.0224 | -104.4856
16.0261 | -104.4741
16.0298 | -104.4653
16.0335 | -104.4569
16.0373 | -104.4482
16.041 | -104.4422
16.0447 | -104.4381
16.0485 | -104.4347
16.0522 | -104.4323
16.0559 | -104.4322
16.0596 | -104.4377
16.0634 | -104.4426
16.0671 | -104.4491
16.0708 | -104.4589
16.0745 | -104.4729
16.0783 | -104.4883
16.082 | -104.5013
16.0857 | -104.5151
16.0894 | -104.5329
16.0932 | -104.5508
16.0969 | -104.5686
16.1006 | -104.5894
16.1044 | -104.6084
16.1081 | -104.6264
16.1118 | -104.6449
16.1155 | -104.6631
16.1193 | -104.6798
16.123 | -104.6979
16.1267 | -104.7179
16.1304 | -104.7278
16.1342 | -104.7429
16.1379 | -104.758
16.1416 | -104.7735
16.1454 | -104.7869
16.1491 | -104.7973
16.1528 | -104.8062
16.1565 | -104.8134
16.1603 | -104.8166
16.164 | -104.8202
16.1677 | -104.8239
16.1714 | -104.8247
16.1752 | -104.824
16.1789 | -104.8203
16.1826 | -104.8174
16.1864 | -104.8196
16.1901 | -104.8221
16.1938 | -104.8189
16.1975 | -104.8122
16.2013 | -104.8056
16.205 | -104.8023
16.2087 | -104.8012
16.2124 | -104.8008
16.2162 | -104.8023
16.2199 | -104.8054
16.2236 | -104.8105
16.2273 | -104.8146
16.2311 | -104.8175
16.2348 | -104.8209
16.2385 | -104.8247
16.2423 | -104.8307
16.246 | -104.8382
16.2497 | -104.8475
16.2534 | -104.8589
16.2572 | -104.8686
16.2609 | -104.878
16.2646 | -104.8872
16.2683 | -104.8974
16.2721 | -104.9059
16.2758 | -104.9145
16.2795 | -104.9246
16.2833 | -104.934
16.287 | -104.9417
16.2907 | -104.95
16.2944 | -104.957
16.2982 | -104.9593
16.3019 | -104.9615
16.3056 | -104.9641
16.3093 | -104.9637
16.3131 | -104.9619
16.3168 | -104.9643
16.3205 | -104.9664
16.3242 | -104.9667
16.328 | -104.9674
16.3317 | -104.9653
16.3354 | -104.9644
16.3392 | -104.9633
16.3429 | -104.9597
16.3466 | -104.9531
16.3503 | -104.9481
16.3541 | -104.9428
16.3578 | -104.9367
16.3615 | -104.9305
16.3652 | -104.9271
16.369 | -104.9207
16.3727 | -104.9113
16.3764 | -104.9011
16.3802 | -104.8925
16.3839 | -104.8833
16.3876 | -104.8753
16.3913 | -104.8659
16.3951 | -104.8569
16.3988 | -104.8464
16.4025 | -104.8362
16.4062 | -104.8271
16.41 | -104.8172
16.4137 | -104.8097
16.4174 | -104.7985
16.4212 | -104.7872
16.4249 | -104.7764
16.4286 | -104.7699
16.4323 | -104.7644
16.4361 | -104.7591
16.4398 | -104.7498
16.4435 | -104.7391
16.4472 | -104.7333
16.451 | -104.7263
16.4547 | -104.7195
16.4584 | -104.7139
16.4621 | -104.7111
16.4659 | -104.707
16.4696 | -104.7043
16.4733 | -104.7034
16.4771 | -104.7007
16.4808 | -104.698
16.4845 | -104.6964
16.4882 | -104.6957
16.492 | -104.6986
16.4957 | -104.7031
16.4994 | -104.7074
16.5031 | -104.7107
16.5069 | -104.7166
16.5106 | -104.7233
16.5143 | -104.7293
16.5181 | -104.7366
16.5218 | -104.7463
16.5255 | -104.7538
16.5292 | -104.7648
16.533 | -104.7808
16.5367 | -104.7947
16.5404 | -104.8056
16.5441 | -104.8208
16.5479 | -104.8398
16.5516 | -104.8562
16.5553 | -104.8705
16.559 | -104.8874
16.5628 | -104.9029
16.5665 | -104.921
16.5702 | -104.943
16.574 | -104.9637
16.5777 | -104.9846
16.5814 | -105.0078
16.5851 | -105.0295
16.5889 | -105.047
16.5926 | -105.0637
16.5963 | -105.0812
16.6 | -105.097
16.6038 | -105.1129
16.6075 | -105.1273
16.6112 | -105.142
16.615 | -105.1536
16.6187 | -105.1623
16.6224 | -105.1679
16.6261 | -105.1686
16.6299 | -105.1681
16.6336 | -105.1683
16.6373 | -105.1652
16.641 | -105.1618
16.6448 | -105.1573
16.6485 | -105.1486
16.6522 | -105.1432
16.656 | -105.1347
16.6597 | -105.1259
16.6634 | -105.114
16.6671 | -105.0976
16.6709 | -105.0861
16.6746 | -105.0766
16.6783 | -105.0671
16.682 | -105.0583
16.6858 | -105.0462
16.6895 | -105.0364
16.6932 | -105.0241
16.6969 | -105.0083
16.7007 | -104.9939
16.7044 | -104.9843
16.7081 | -104.9744
16.7119 | -104.9616
16.7156 | -104.9475
16.7193 | -104.9318
16.723 | -104.9161
16.7268 | -104.9001
16.7305 | -104.8848
16.7342 | -104.8704
16.7379 | -104.8546
16.7417 | -104.8387
16.7454 | -104.8213
16.7491 | -104.8041
16.7529 | -104.7874
16.7566 | -104.7723
16.7603 | -104.7563
16.764 | -104.7387
16.7678 | -104.7195
16.7715 | -104.7002
16.7752 | -104.6836
16.7789 | -104.6669
16.7827 | -104.6496
16.7864 | -104.6331
16.7901 | -104.6156
16.7938 | -104.6
16.7976 | -104.5849
16.8013 | -104.5724
16.805 | -104.5615
16.8088 | -104.5495
16.8125 | -104.5414
16.8162 | -104.5339
16.8199 | -104.5279
16.8237 | -104.526
16.8274 | -104.5287
16.8311 | -104.5342
16.8348 | -104.5398
16.8386 | -104.5468
16.8423 | -104.5579
16.846 | -104.5728
16.8498 | -104.5874
16.8535 | -104.6049
16.8572 | -104.6259
16.8609 | -104.6474
16.8647 | -104.6737
16.8684 | -104.6994
16.8721 | -104.7245
16.8758 | -104.7518
16.8796 | -104.779
16.8833 | -104.8069
16.887 | -104.8375
16.8908 | -104.8684
16.8945 | -104.8995
16.8982 | -104.929
16.9019 | -104.9596
16.9057 | -104.9914
16.9094 | -105.0202
16.9131 | -105.0471
16.9168 | -105.0715
16.9206 | -105.0969
16.9243 | -105.123
16.928 | -105.1475
16.9317 | -105.1671
16.9355 | -105.1874
16.9392 | -105.2082
16.9429 | -105.2245
16.9467 | -105.2387
16.9504 | -105.252
16.9541 | -105.2646
16.9578 | -105.2747
16.9616 | -105.2868
16.9653 | -105.2935
16.969 | -105.297
16.9727 | -105.2983
16.9765 | -105.2988
16.9802 | -105.3011
16.9839 | -105.3026
16.9877 | -105.3014
16.9914 | -105.297
16.9951 | -105.2929
16.9988 | -105.2888
17.0026 | -105.2858
17.0063 | -105.2842
17.01 | -105.2806
17.0137 | -105.2766
17.0175 | -105.2722
17.0212 | -105.2696
17.0249 | -105.2672
17.0287 | -105.2612
17.0324 | -105.2614
17.0361 | -105.2615
17.0398 | -105.2602
17.0436 | -105.2591
17.0473 | -105.2571
17.051 | -105.2561
17.0547 | -105.2544
17.0585 | -105.2545
17.0622 | -105.2517
17.0659 | -105.2483
17.0696 | -105.2468
17.0734 | -105.2429
17.0771 | -105.2414
17.0808 | -105.2403
17.0846 | -105.238
17.0883 | -105.2357
17.092 | -105.2335
17.0957 | -105.2319
17.0995 | -105.23
17.1032 | -105.2293
17.1069 | -105.2263
17.1106 | -105.2241
17.1144 | -105.2219
17.1181 | -105.2188
17.1218 | -105.2176
17.1256 | -105.2165
17.1293 | -105.2129
17.133 | -105.2069
17.1367 | -105.2018
17.1405 | -105.1987
17.1442 | -105.1945
17.1479 | -105.1916
17.1516 | -105.1888
17.1554 | -105.1863
17.1591 | -105.1833
17.1628 | -105.1786
17.1666 | -105.172
17.1703 | -105.1655
17.174 | -105.1611
17.1777 | -105.1555
17.1815 | -105.1493
17.1852 | -105.1451
17.1889 | -105.1379
17.1926 | -105.1298
17.1964 | -105.1238
17.2001 | -105.1196
17.2038 | -105.1128
17.2075 | -105.106
17.2113 | -105.1008
17.215 | -105.0928
17.2187 | -105.0829
17.2225 | -105.0733
17.2262 | -105.0641
17.2299 | -105.0562
17.2336 | -105.0497
17.2374 | -105.0443
17.2411 | -105.0368
17.2448 | -105.0289
17.2485 | -105.0207
17.2523 | -105.0104
17.256 | -105.0031
17.2597 | -104.9972
17.2635 | -104.9923
17.2672 | -104.99
17.2709 | -104.9865
17.2746 | -104.9834
17.2784 | -104.9795
17.2821 | -104.9766
17.2858 | -104.9744
17.2895 | -104.9764
17.2933 | -104.9783
17.297 | -104.9808
17.3007 | -104.9841
17.3044 | -104.9858
17.3082 | -104.9931
17.3119 | -105.0009
17.3156 | -105.006
17.3194 | -105.0149
17.3231 | -105.024
17.3268 | -105.0325
17.3305 | -105.0409
17.3343 | -105.0484
17.338 | -105.0595
17.3417 | -105.0705
17.3454 | -105.0818
17.3492 | -105.0935
17.3529 | -105.1065
17.3566 | -105.1175
17.3604 | -105.1307
17.3641 | -105.1437
17.3678 | -105.1538
17.3715 | -105.1648
17.3753 | -105.1753
17.379 | -105.1879
17.3827 | -105.1978
17.3864 | -105.2077
17.3902 | -105.2182
17.3939 | -105.2251
17.3976 | -105.2332
17.4014 | -105.243
17.4051 | -105.2512
17.4088 | -105.2579
17.4125 | -105.2626
17.4163 | -105.2707
17.42 | -105.2767
17.4237 | -105.2839
17.4274 | -105.2903
17.4312 | -105.2926
17.4349 | -105.2951
17.4386 | -105.2963
17.4423 | -105.2971
17.4461 | -105.2977
17.4498 | -105.2973
17.4535 | -105.2965
17.4573 | -105.2991
17.461 | -105.302
17.4647 | -105.3054
17.4684 | -105.3074
17.4722 | -105.3081
17.4759 | -105.308
17.4796 | -105.3064
17.4833 | -105.3045
17.4871 | -105.2998
17.4908 | -105.2962
17.4945 | -105.2936
17.4983 | -105.2906
17.502 | -105.2902
17.5057 | -105.2891
17.5094 | -105.2838
17.5132 | -105.2775
17.5169 | -105.2714
17.5206 | -105.2653
17.5243 | -105.2619
17.5281 | -105.2566
17.5318 | -105.2506
17.5355 | -105.246
17.5392 | -105.2363
17.543 | -105.2293
17.5467 | -105.2216
17.5504 | -105.215
17.5542 | -105.2095
17.5579 | -105.2006
17.5616 | -105.1894
17.5653 | -105.1811
17.5691 | -105.1709
17.5728 | -105.1575
17.5765 | -105.1446
17.5802 | -105.1309
17.584 | -105.1187
17.5877 | -105.1059
17.5914 | -105.0906
17.5952 | -105.075
17.5989 | -105.0611
17.6026 | -105.0484
17.6063 | -105.0368
17.6101 | -105.0196
17.6138 | -105.0041
17.6175 | -104.99
17.6212 | -104.9743
17.625 | -104.9618
17.6287 | -104.9467
17.6324 | -104.9319
17.6362 | -104.9165
17.6399 | -104.9017
17.6436 | -104.8889
17.6473 | -104.8732
17.6511 | -104.8568
17.6548 | -104.8437
17.6585 | -104.8317
17.6622 | -104.8205
17.666 | -104.8062
17.6697 | -104.7921
17.6734 | -104.7787
17.6771 | -104.7682
17.6809 | -104.7598
17.6846 | -104.7533
17.6883 | -104.7482
17.6921 | -104.7427
17.6958 | -104.7401
17.6995 | -104.7378
17.7032 | -104.734
17.707 | -104.7279
17.7107 | -104.7235
17.7144 | -104.7203
17.7181 | -104.7185
17.7219 | -104.7204
17.7256 | -104.7242
17.7293 | -104.7278
17.7331 | -104.7312
17.7368 | -104.736
17.7405 | -104.7387
17.7442 | -104.7432
17.748 | -104.7461
17.7517 | -104.7495
17.7554 | -104.7529
17.7591 | -104.7542
17.7629 | -104.7576
17.7666 | -104.761
17.7703 | -104.765
17.7741 | -104.7702
17.7778 | -104.7723
17.7815 | -104.7749
17.7852 | -104.7777
17.789 | -104.7816
17.7927 | -104.7864
17.7964 | -104.7889
17.8001 | -104.7926
17.8039 | -104.7972
17.8076 | -104.8019
17.8113 | -104.8077
17.815 | -104.8151
17.8188 | -104.8214
17.8225 | -104.8282
17.8262 | -104.8376
17.83 | -104.8474
17.8337 | -104.8585
17.8374 | -104.8676
17.8411 | -104.8736
17.8449 | -104.8816
17.8486 | -104.8881
17.8523 | -104.8939
17.856 | -104.8993
17.8598 | -104.907
17.8635 | -104.9169
17.8672 | -104.9255
17.871 | -104.9335
17.8747 | -104.9439
17.8784 | -104.9519
17.8821 | -104.9605
17.8859 | -104.9678
17.8896 | -104.9737
17.8933 | -104.9797
17.897 | -104.9859
17.9008 | -104.9928
17.9045 | -104.9984
17.9082 | -105.0032
17.912 | -105.0076
17.9157 | -105.0131
17.9194 | -105.0171
17.9231 | -105.0228
17.9269 | -105.0291
17.9306 | -105.0317
17.9343 | -105.0313
17.938 | -105.0269
17.9418 | -105.022
17.9455 | -105.0185
17.9492 | -105.0154
17.9529 | -105.0131
17.9567 | -105.0114
17.9604 | -105.0087
17.9641 | -105.0051
17.9679 | -105.0015
17.9716 | -104.9941
17.9753 | -104.9831
17.979 | -104.9726
17.9828 | -104.9621
17.9865 | -104.9517
17.9902 | -104.9418
17.9939 | -104.932
17.9977 | -104.9194
18.0014 | -104.9075
18.0051 | -104.8943
18.0089 | -104.8804
18.0126 | -104.8681
18.0163 | -104.8584
18.02 | -104.8479
18.0238 | -104.8364
18.0275 | -104.8242
18.0312 | -104.8137
18.0349 | -104.8072
18.0387 | -104.8006
18.0424 | -104.7956
18.0461 | -104.7873
18.0498 | -104.7779
18.0536 | -104.7723
18.0573 | -104.7697
18.061 | -104.7679
18.0648 | -104.7658
18.0685 | -104.7648
18.0722 | -104.7668
18.0759 | -104.7697
18.0797 | -104.7736
18.0834 | -104.7774
18.0871 | -104.7813
18.0908 | -104.7858
18.0946 | -104.7929
18.0983 | -104.8013
18.102 | -104.811
18.1058 | -104.8229
18.1095 | -104.8341
18.1132 | -104.8451
18.1169 | -104.8554
18.1207 | -104.8661
18.1244 | -104.8787
18.1281 | -104.892
18.1318 | -104.9023
18.1356 | -104.9127
18.1393 | -104.9263
18.143 | -104.9404
18.1468 | -104.9519
18.1505 | -104.9633
18.1542 | -104.9746
18.1579 | -104.9849
18.1617 | -104.996
18.1654 | -105.0074
18.1691 | -105.0183
18.1728 | -105.0276
18.1766 | -105.037
18.1803 | -105.0449
18.184 | -105.0513
18.1877 | -105.0601
18.1915 | -105.0698
18.1952 | -105.0807
18.1989 | -105.0899
18.2027 | -105.0963
18.2064 | -105.0991
18.2101 | -105.1024
18.2138 | -105.1064
18.2176 | -105.1094
18.2213 | -105.1115
18.225 | -105.1131
18.2287 | -105.1153
18.2325 | -105.1151
18.2362 | -105.1122
18.2399 | -105.1093
18.2437 | -105.1059
18.2474 | -105.1029
18.2511 | -105.1017
18.2548 | -105.0979
18.2586 | -105.0922
18.2623 | -105.0869
18.266 | -105.0786
18.2697 | -105.068
18.2735 | -105.0602
18.2772 | -105.0527
18.2809 | -105.0453
18.2846 | -105.0372
18.2884 | -105.0271
18.2921 | -105.0168
18.2958 | -105.0071
18.2996 | -104.9961
18.3033 | -104.984
18.307 | -104.9774
18.3107 | -104.9746
18.3145 | -104.9688
18.3182 | -104.9598
18.3219 | -104.9507
18.3256 | -104.9429
18.3294 | -104.9364
18.3331 | -104.9277
18.3368 | -104.9203
18.3406 | -104.9144
18.3443 | -104.9084
18.348 | -104.9012
18.3517 | -104.8924
18.3555 | -104.8843
18.3592 | -104.8776
18.3629 | -104.8713
18.3666 | -104.8641
18.3704 | -104.8583
18.3741 | -104.8531
18.3778 | -104.8466
18.3816 | -104.8419
18.3853 | -104.8362
18.389 | -104.8298
18.3927 | -104.8234
18.3965 | -104.8181
18.4002 | -104.8111
18.4039 | -104.8069
18.4076 | -104.8044
18.4114 | -104.8005
18.4151 | -104.7925
18.4188 | -104.7874
18.4225 | -104.7825
18.4263 | -104.7785
18.43 | -104.7762
18.4337 | -104.7733
18.4375 | -104.7714
18.4412 | -104.7686
18.4449 | -104.7651
18.4486 | -104.7599
18.4524 | -104.7546
18.4561 | -104.7517
18.4598 | -104.7492
18.4635 | -104.7461
18.4673 | -104.7436
18.471 | -104.7389
18.4747 | -104.7356
18.4785 | -104.7308
18.4822 | -104.7258
18.4859 | -104.7218
18.4896 | -104.7164
18.4934 | -104.7127
18.4971 | -104.7106
18.5008 | -104.7089
18.5045 | -104.7062
18.5083 | -104.7016
18.512 | -104.6964
18.5157 | -104.694
18.5194 | -104.6921
18.5232 | -104.688
18.5269 | -104.6848
18.5306 | -104.6823
18.5344 | -104.6752
18.5381 | -104.6717
18.5418 | -104.6691
18.5455 | -104.6657
18.5493 | -104.6655
18.553 | -104.6659
18.5567 | -104.6666
18.5604 | -104.6662
18.5642 | -104.6672
18.5679 | -104.6722
18.5716 | -104.6749
18.5754 | -104.6764
18.5791 | -104.679
18.5828 | -104.6832
18.5865 | -104.6888
18.5903 | -104.6957
18.594 | -104.7018
18.5977 | -104.7089
18.6014 | -104.717
18.6052 | -104.7245
18.6089 | -104.7331
18.6126 | -104.7462
18.6164 | -104.7606
18.6201 | -104.7732
18.6238 | -104.7842
18.6275 | -104.7949
18.6313 | -104.8045
18.635 | -104.8134
18.6387 | -104.8238
18.6424 | -104.8342
18.6462 | -104.8416
18.6499 | -104.8489
18.6536 | -104.8583
18.6573 | -104.8693
18.6611 | -104.8796
18.6648 | -104.8872
18.6685 | -104.8935
18.6723 | -104.9006
18.676 | -104.9083
18.6797 | -104.9172
18.6834 | -104.9221
18.6872 | -104.9262
18.6909 | -104.9268
18.6946 | -104.9283
18.6983 | -104.9301
18.7021 | -104.9301
18.7058 | -104.9311
18.7095 | -104.9339
18.7133 | -104.9352
18.717 | -104.9366
18.7207 | -104.9353
18.7244 | -104.936
18.7282 | -104.9375
18.7319 | -104.939
18.7356 | -104.9406
18.7393 | -104.9404
18.7431 | -104.9409
18.7468 | -104.9411
18.7505 | -104.9384
18.7543 | -104.9394
18.758 | -104.9401
18.7617 | -104.9411
18.7654 | -104.9399
18.7692 | -104.9384
18.7729 | -104.9378
18.7766 | -104.9363
18.7803 | -104.9344
18.7841 | -104.9334
18.7878 | -104.9355
18.7915 | -104.9353
18.7952 | -104.9333
18.799 | -104.9328
18.8027 | -104.9316
18.8064 | -104.931
18.8102 | -104.9298
18.8139 | -104.9269
18.8176 | -104.9254
18.8213 | -104.9245
18.8251 | -104.9247
18.8288 | -104.9259
18.8325 | -104.9265
18.8362 | -104.9249
18.84 | -104.9244
18.8437 | -104.9242
18.8474 | -104.9238
18.8512 | -104.9222
18.8549 | -104.9215
18.8586 | -104.9232
18.8623 | -104.9251
18.8661 | -104.9253
18.8698 | -104.9261
18.8735 | -104.9255
18.8772 | -104.9258
18.881 | -104.9252
18.8847 | -104.9262
18.8884 | -104.9264
18.8922 | -104.9276
18.8959 | -104.9296
18.8996 | -104.9318
18.9033 | -104.9333
18.9071 | -104.9344
18.9108 | -104.9363
18.9145 | -104.9393
18.9182 | -104.9417
18.922 | -104.943
18.9257 | -104.9422
18.9294 | -104.9398
18.9331 | -104.9397
18.9369 | -104.9405
18.9406 | -104.9388
18.9443 | -104.9386
18.9481 | -104.9403
18.9518 | -104.9411
18.9555 | -104.9418
18.9592 | -104.943
18.963 | -104.942
18.9667 | -104.9388
18.9704 | -104.938
18.9741 | -104.9385
18.9779 | -104.9384
18.9816 | -104.9376
18.9853 | -104.9371
18.9891 | -104.9375
18.9928 | -104.9343
18.9965 | -104.9296
19.0002 | -104.9285
19.004 | -104.9305
19.0077 | -104.9306
19.0114 | -104.9307
19.0151 | -104.9323
19.0189 | -104.9345
19.0226 | -104.9341
19.0263 | -104.9329
19.03 | -104.9302
19.0338 | -104.9292
19.0375 | -104.929
19.0412 | -104.9281
19.045 | -104.9249
19.0487 | -104.9234
19.0524 | -104.9212
19.0561 | -104.9169
19.0599 | -104.9166
19.0636 | -104.9158
19.0673 | -104.9115
19.071 | -104.9048
19.0748 | -104.8977
19.0785 | -104.894
19.0822 | -104.8916
19.086 | -104.8881
19.0897 | -104.8838
19.0934 | -104.8774
19.0971 | -104.8706
19.1009 | -104.8655
19.1046 | -104.8616
19.1083 | -104.8563
19.112 | -104.8514
19.1158 | -104.8472
19.1195 | -104.8432
19.1232 | -104.8398
19.127 | -104.8325
19.1307 | -104.8208
19.1344 | -104.8109
19.1381 | -104.8029
19.1419 | -104.7936
19.1456 | -104.7868
19.1493 | -104.7789
19.153 | -104.7685
19.1568 | -104.7588
19.1605 | -104.75
19.1642 | -104.7405
19.1679 | -104.7311
19.1717 | -104.7225
19.1754 | -104.712
19.1791 | -104.6995
19.1829 | -104.6873
19.1866 | -104.6732
19.1903 | -104.6593
19.194 | -104.6455
19.1978 | -104.634
19.2015 | -104.6234
19.2052 | -104.6115
19.2089 | -104.5988
19.2127 | -104.5871
19.2164 | -104.5721
19.2201 | -104.5571
19.2239 | -104.5438
19.2276 | -104.5305
19.2313 | -104.5177
19.235 | -104.5077
19.2388 | -104.4991
19.2425 | -104.4867
19.2462 | -104.473
19.2499 | -104.4605
19.2537 | -104.4513
19.2574 | -104.443
19.2611 | -104.4357
19.2648 | -104.4293
19.2686 | -104.4239
19.2723 | -104.4174
19.276 | -104.4093
19.2798 | -104.4036
19.2835 | -104.3996
19.2872 | -104.3953
19.2909 | -104.3925
19.2947 | -104.391
19.2984 | -104.3905
19.3021 | -104.3901
19.3058 | -104.3907
19.3096 | -104.3909
19.3133 | -104.3897
19.317 | -104.3903
19.3208 | -104.3907
19.3245 | -104.3904
19.3282 | -104.391
19.3319 | -104.3917
19.3357 | -104.3923
19.3394 | -104.3922
19.3431 | -104.3923
19.3468 | -104.3922
19.3506 | -104.3891
19.3543 | -104.3881
19.358 | -104.3869
19.3618 | -104.3844
19.3655 | -104.3832
19.3692 | -104.3833
19.3729 | -104.3837
19.3767 | -104.3828
19.3804 | -104.3822
19.3841 | -104.3818
19.3878 | -104.3822
19.3916 | -104.3817
19.3953 | -104.3826
19.399 | -104.3832
19.4027 | -104.3835
19.4065 | -104.3854
19.4102 | -104.3861
19.4139 | -104.3861
19.4177 | -104.387
19.4214 | -104.3874
19.4251 | -104.3881
19.4288 | -104.3889
19.4326 | -104.388
19.4363 | -104.3873
19.44 | -104.3878
19.4437 | -104.3889
19.4475 | -104.3907
19.4512 | -104.3947
19.4549 | -104.3989
19.4587 | -104.4028
19.4624 | -104.4076
19.4661 | -104.4088
19.4698 | -104.4105
19.4736 | -104.4119
19.4773 | -104.411
19.481 | -104.4109
19.4847 | -104.4113
19.4885 | -104.4115
19.4922 | -104.4136
19.4959 | -104.4141
19.4997 | -104.4128
19.5034 | -104.4111
19.5071 | -104.409
19.5108 | -104.4041
19.5146 | -104.3996
19.5183 | -104.3972
19.522 | -104.3948
19.5257 | -104.3904
19.5295 | -104.3862
19.5332 | -104.3796
19.5369 | -104.3714
19.5406 | -104.3665
19.5444 | -104.3613
19.5481 | -104.3544
19.5518 | -104.3496
19.5556 | -104.3434
19.5593 | -104.3358
19.563 | -104.3301
19.5667 | -104.324
19.5705 | -104.3172
19.5742 | -104.3102
19.5779 | -104.3048
19.5816 | -104.2986
19.5854 | -104.2937
19.5891 | -104.2931
19.5928 | -104.2907
19.5966 | -104.2859
19.6003 | -104.2823
19.604 | -104.2785
19.6077 | -104.2729
19.6115 | -104.2692
19.6152 | -104.2677
19.6189 | -104.2667
19.6226 | -104.2651
19.6264 | -104.266
19.6301 | -104.2656
19.6338 | -104.262
19.6375 | -104.2615
19.6413 | -104.2628
19.645 | -104.2629
19.6487 | -104.2615
19.6525 | -104.2611
19.6562 | -104.2628
19.6599 | -104.2648
19.6636 | -104.2646
19.6674 | -104.2616
19.6711 | -104.2605
19.6748 | -104.2542
19.6785 | -104.2515
19.6823 | -104.2487
19.686 | -104.2434
19.6897 | -104.2391
19.6935 | -104.2329
19.6972 | -104.2245
19.7009 | -104.2155
19.7046 | -104.2041
19.7084 | -104.1915
19.7121 | -104.1788
19.7158 | -104.1654
19.7195 | -104.1521
19.7233 | -104.1358
19.727 | -104.1188
19.7307 | -104.0992
19.7345 | -104.0795
19.7382 | -104.0594
19.7419 | -104.0391
19.7456 | -104.0208
19.7494 | -104.0008
19.7531 | -103.9785
19.7568 | -103.9583
19.7605 | -103.9376
19.7643 | -103.9155
19.768 | -103.8928
19.7717 | -103.87
19.7754 | -103.8491
19.7792 | -103.8279
19.7829 | -103.8084
19.7866 | -103.7909
19.7904 | -103.7746
19.7941 | -103.7579
19.7978 | -103.7424
19.8015 | -103.7261
19.8053 | -103.7115
19.809 | -103.6979
19.8127 | -103.6857
19.8164 | -103.6752
19.8202 | -103.6646
19.8239 | -103.6568
19.8276 | -103.6516
19.8314 | -103.6462
19.8351 | -103.6439
19.8388 | -103.6434
19.8425 | -103.6423
19.8463 | -103.6419
19.85 | -103.6435
19.8537 | -103.6464
19.8574 | -103.6478
19.8612 | -103.6509
19.8649 | -103.6559
19.8686 | -103.6608
19.8724 | -103.6659
19.8761 | -103.672
19.8798 | -103.6802
19.8835 | -103.6881
19.8873 | -103.6959
19.891 | -103.7022
19.8947 | -103.7099
19.8984 | -103.7176
19.9022 | -103.7236
19.9059 | -103.7292
19.9096 | -103.7349
19.9133 | -103.7425
19.9171 | -103.7524
19.9208 | -103.7592
19.9245 | -103.766
19.9283 | -103.7739
19.932 | -103.7818
19.9357 | -103.7871
19.9394 | -103.7904
19.9432 | -103.7945
19.9469 | -103.8021
19.9506 | -103.8097
19.9543 | -103.8155
19.9581 | -103.8246
19.9618 | -103.8286
19.9655 | -103.8321
19.9693 | -103.8367
19.973 | -103.8374
19.9767 | -103.8394
19.9804 | -103.8408
19.9842 | -103.8405
19.9879 | -103.8423
19.9916 | -103.8437
19.9953 | -103.8438
19.9991 | -103.8428
20.0028 | -103.8394
20.0065 | -103.8361
20.0102 | -103.832
20.014 | -103.8272
20.0177 | -103.8186
20.0214 | -103.8075
20.0252 | -103.7977
20.0289 | -103.7854
20.0326 | -103.7718
20.0363 | -103.7595
20.0401 | -103.7447
20.0438 | -103.728
20.0475 | -103.713
20.0512 | -103.6986
20.055 | -103.6821
20.0587 | -103.6633
20.0624 | -103.6452
20.0662 | -103.6257
20.0699 | -103.6038
20.0736 | -103.5805
20.0773 | -103.5602
20.0811 | -103.54
20.0848 | -103.5212
20.0885 | -103.5013
20.0922 | -103.4798
20.096 | -103.4628
20.0997 | -103.4463
20.1034 | -103.4298
20.1072 | -103.4136
20.1109 | -103.398
20.1146 | -103.3846
20.1183 | -103.3721
20.1221 | -103.3612
20.1258 | -103.3528
20.1295 | -103.3485
20.1332 | -103.3457
20.137 | -103.3431
20.1407 | -103.345
20.1444 | -103.3478
20.1481 | -103.3494
20.1519 | -103.3524
20.1556 | -103.36
20.1593 | -103.3674
20.1631 | -103.3748
20.1668 | -103.3851
20.1705 | -103.3971
20.1742 | -103.41
20.178 | -103.4231
20.1817 | -103.4352
20.1854 | -103.448
20.1891 | -103.4609
20.1929 | -103.4732
20.1966 | -103.485
20.2003 | -103.4961
20.2041 | -103.5067
20.2078 | -103.5164
20.2115 | -103.5245
20.2152 | -103.532
20.219 | -103.5414
20.2227 | -103.5483
20.2264 | -103.5523
20.2301 | -103.5555
20.2339 | -103.5561
20.2376 | -103.5547
20.2413 | -103.5519
20.245 | -103.5466
20.2488 | -103.5403
20.2525 | -103.5329
20.2562 | -103.5244
20.26 | -103.5163
20.2637 | -103.5068
20.2674 | -103.4962
20.2711 | -103.4855
20.2749 | -103.4729
20.2786 | -103.4603
20.2823 | -103.4477
20.286 | -103.4343
20.2898 | -103.4203
20.2935 | -103.4073
20.2972 | -103.3945
20.301 | -103.3843
20.3047 | -103.3758
20.3084 | -103.3677
20.3121 | -103.3585
20.3159 | -103.349
20.3196 | -103.3393
20.3233 | -103.3306
20.327 | -103.324
20.3308 | -103.3188
20.3345 | -103.3111
20.3382 | -103.3049
20.342 | -103.3007
20.3457 | -103.2971
20.3494 | -103.2911
20.3531 | -103.2825
20.3569 | -103.2741
20.3606 | -103.2687
20.3643 | -103.262
20.368 | -103.2531
20.3718 | -103.2437
20.3755 | -103.2341
20.3792 | -103.2253
20.3829 | -103.2135
20.3867 | -103.2004
20.3904 | -103.1868
20.3941 | -103.1708
20.3979 | -103.1534
20.4016 | -103.1348
20.4053 | -103.1145
20.409 | -103.0918
20.4128 | -103.0674
20.4165 | -103.041
20.4202 | -103.0139
20.4239 | -102.9876
20.4277 | -102.96
20.4314 | -102.9361
20.4351 | -102.9113
20.4389 | -102.8887
20.4426 | -102.8672
20.4463 | -102.8456
20.45 | -102.8272
20.4538 | -102.8102
20.4575 | -102.7932
20.4612 | -102.7782
20.4649 | -102.7662
20.4687 | -102.7554
20.4724 | -102.7475
20.4761 | -102.7424
20.4799 | -102.7409
20.4836 | -102.7416
20.4873 | -102.7441
20.491 | -102.7469
20.4948 | -102.751
20.4985 | -102.7588
20.5022 | -102.7692
20.5059 | -102.7814
20.5097 | -102.7932
20.5134 | -102.8073
20.5171 | -102.824
20.5208 | -102.8404
20.5246 | -102.8578
20.5283 | -102.8768
20.532 | -102.8963
20.5358 | -102.9164
20.5395 | -102.938
20.5432 | -102.9585
20.5469 | -102.9771
20.5507 | -102.995
20.5544 | -103.0133
20.5581 | -103.0319
20.5618 | -103.0501
20.5656 | -103.0677
20.5693 | -103.0843
20.573 | -103.0995
20.5768 | -103.1146
20.5805 | -103.1286
20.5842 | -103.1421
20.5879 | -103.1552
20.5917 | -103.1664
20.5954 | -103.1776
20.5991 | -103.1896
20.6028 | -103.2007
20.6066 | -103.2115
20.6103 | -103.222
20.614 | -103.2294
20.6178 | -103.2372
20.6215 | -103.2449
20.6252 | -103.2527
20.6289 | -103.2585
20.6327 | -103.2645
20.6364 | -103.2707
20.6401 | -103.2771
20.6438 | -103.2854
20.6476 | -103.2928
20.6513 | -103.2976
20.655 | -103.302
20.6587 | -103.3068
20.6625 | -103.3117
20.6662 | -103.3137
20.6699 | -103.312
20.6737 | -103.3106
20.6774 | -103.3102
20.6811 | -103.3091
20.6848 | -103.3084
20.6886 | -103.3072
20.6923 | -103.3059
20.696 | -103.3035
20.6997 | -103.3006
20.7035 | -103.2957
20.7072 | -103.2899
20.7109 | -103.2821
20.7147 | -103.2719
20.7184 | -103.2609
20.7221 | -103.2487
20.7258 | -103.2369
20.7296 | -103.2251
20.7333 | -103.2121
20.737 | -103.1967
20.7407 | -103.1808
20.7445 | -103.1633
20.7482 | -103.144
20.7519 | -103.1245
20.7556 | -103.1062
20.7594 | -103.0866
20.7631 | -103.0659
20.7668 | -103.0473
20.7706 | -103.0283
20.7743 | -103.0096
20.778 | -102.9927
20.7817 | -102.9751
20.7855 | -102.9581
20.7892 | -102.9414
20.7929 | -102.9256
20.7966 | -102.9109
20.8004 | -102.8953
20.8041 | -102.8812
20.8078 | -102.871
20.8116 | -102.8619
20.8153 | -102.8533
20.819 | -102.845
20.8227 | -102.8368
20.8265 | -102.8312
20.8302 | -102.8278
20.8339 | -102.822
20.8376 | -102.817
20.8414 | -102.8133
20.8451 | -102.8076
20.8488 | -102.8015
20.8526 | -102.7965
20.8563 | -102.7915
20.86 | -102.7853
20.8637 | -102.7803
20.8675 | -102.7742
20.8712 | -102.7651
20.8749 | -102.7558
20.8786 | -102.7446
20.8824 | -102.7329
20.8861 | -102.7198
20.8898 | -102.7035
20.8935 | -102.6863
20.8973 | -102.6689
20.901 | -102.6513
20.9047 | -102.6316
20.9085 | -102.6123
20.9122 | -102.5897
20.9159 | -102.5649
20.9196 | -102.5404
20.9234 | -102.516
20.9271 | -102.4904
20.9308 | -102.4643
20.9345 | -102.4391
20.9383 | -102.4112
20.942 | -102.3831
20.9457 | -102.3544
20.9495 | -102.3269
20.9532 | -102.3016
20.9569 | -102.2767
20.9606 | -102.2548
20.9644 | -102.2354
20.9681 | -102.2275
20.9718 | -102.2113
20.9755 | -102.197
20.9793 | -102.1845
20.983 | -102.1717
20.9867 | -102.1602
20.9904 | -102.1513
20.9942 | -102.1456
20.9979 | -102.1428
21.0016 | -102.1423
21.0054 | -102.1441
21.0091 | -102.1482
21.0128 | -102.1556
21.0165 | -102.1648
21.0203 | -102.176
21.024 | -102.1893
21.0277 | -102.2043
21.0314 | -102.2214
21.0352 | -102.2412
21.0389 | -102.2614
21.0426 | -102.2848
21.0464 | -102.3113
21.0501 | -102.3398
21.0538 | -102.3675
21.0575 | -102.3965
21.0613 | -102.4265
21.065 | -102.455
21.0687 | -102.4836
21.0724 | -102.5113
21.0762 | -102.5377
21.0799 | -102.5652
21.0836 | -102.5942
21.0874 | -102.6214
21.0911 | -102.6435
21.0948 | -102.6644
21.0985 | -102.6874
21.1023 | -102.7084
21.106 | -102.7249
21.1097 | -102.7396
21.1134 | -102.7538
21.1172 | -102.7653
21.1209 | -102.7763
21.1246 | -102.7835
21.1283 | -102.7884
21.1321 | -102.792
21.1358 | -102.7929
21.1395 | -102.7909
21.1433 | -102.7896
21.147 | -102.7869
21.1507 | -102.7798
21.1544 | -102.7712
21.1582 | -102.7618
21.1619 | -102.7498
21.1656 | -102.737
21.1693 | -102.7245
21.1731 | -102.711
21.1768 | -102.697
21.1805 | -102.6809
21.1843 | -102.6646
21.188 | -102.65
21.1917 | -102.634
21.1954 | -102.617
21.1992 | -102.6011
21.2029 | -102.5866
21.2066 | -102.5721
21.2103 | -102.5575
21.2141 | -102.5438
21.2178 | -102.5303
21.2215 | -102.5182
21.2252 | -102.5064
21.229 | -102.4965
21.2327 | -102.4885
21.2364 | -102.484
21.2402 | -102.4819
21.2439 | -102.4811
21.2476 | -102.4797
21.2513 | -102.4788
21.2551 | -102.4802
21.2588 | -102.484
21.2625 | -102.4894
21.2662 | -102.4951
21.27 | -102.5024
21.2737 | -102.511
21.2774 | -102.5213
21.2812 | -102.5319
21.2849 | -102.5464
21.2886 | -102.5641
21.2923 | -102.5799
21.2961 | -102.5952
21.2998 | -102.6111
21.3035 | -102.6273
21.3072 | -102.6431
21.311 | -102.6606
21.3147 | -102.6777
21.3184 | -102.6947
21.3222 | -102.7105
21.3259 | -102.7265
21.3296 | -102.7445
21.3333 | -102.7622
21.3371 | -102.7787
21.3408 | -102.7945
21.3445 | -102.8085
21.3482 | -102.8213
21.352 | -102.833
21.3557 | -102.8444
21.3594 | -102.8544
21.3631 | -102.8639
21.3669 | -102.8736
21.3706 | -102.8831
21.3743 | -102.8912
21.3781 | -102.8994
21.3818 | -102.906
21.3855 | -102.9101
21.3892 | -102.9144
21.393 | -102.9205
21.3967 | -102.9223
21.4004 | -102.9207
21.4041 | -102.9192
21.4079 | -102.9164
21.4116 | -102.9128
21.4153 | -102.9075
21.4191 | -102.9029
21.4228 | -102.8965
21.4265 | -102.8905
21.4302 | -102.8836
21.434 | -102.8785
21.4377 | -102.8738
21.4414 | -102.869
21.4451 | -102.863
21.4489 | -102.8576
21.4526 | -102.8535
21.4563 | -102.8479
21.4601 | -102.8426
21.4638 | -102.8375
21.4675 | -102.8337
21.4712 | -102.8319
21.475 | -102.8317
21.4787 | -102.8321
21.4824 | -102.8331
21.4861 | -102.8354
21.4899 | -102.8388
21.4936 | -102.8427
21.4973 | -102.8469
21.501 | -102.8516
21.5048 | -102.8641
21.5085 | -102.87
21.5122 | -102.8768
21.516 | -102.8856
21.5197 | -102.8961
21.5234 | -102.9056
21.5271 | -102.9133
21.5309 | -102.9219
21.5346 | -102.9312
21.5383 | -102.9413
21.542 | -102.9504
21.5458 | -102.958
21.5495 | -102.9678
21.5532 | -102.9773
21.557 | -102.9855
21.5607 | -102.9942
21.5644 | -103.0007
21.5681 | -103.0077
21.5719 | -103.0158
21.5756 | -103.0231
21.5793 | -103.0299
21.583 | -103.0355
21.5868 | -103.0428
21.5905 | -103.0495
21.5942 | -103.0534
21.598 | -103.0573
21.6017 | -103.0625
21.6054 | -103.0684
21.6091 | -103.074
21.6129 | -103.0798
21.6166 | -103.0839
21.6203 | -103.0868
21.624 | -103.0904
21.6278 | -103.0948
21.6315 | -103.0977
21.6352 | -103.0997
21.6389 | -103.1015
21.6427 | -103.1049
21.6464 | -103.1099
21.6501 | -103.1123
21.6539 | -103.1138
21.6576 | -103.1165
21.6613 | -103.1177
21.665 | -103.1195
21.6688 | -103.1213
21.6725 | -103.1213
21.6762 | -103.1204
21.6799 | -103.121
21.6837 | -103.1225
21.6874 | -103.1237
21.6911 | -103.1253
21.6949 | -103.1266
21.6986 | -103.1266
21.7023 | -103.1283
21.706 | -103.1285
21.7098 | -103.1293
21.7135 | -103.13
21.7172 | -103.1305
21.7209 | -103.1322
21.7247 | -103.1342
21.7284 | -103.1348
21.7321 | -103.1339
21.7358 | -103.1336
21.7396 | -103.1317
21.7433 | -103.1304
21.747 | -103.1313
21.7508 | -103.1315
21.7545 | -103.1292
21.7582 | -103.1255
21.7619 | -103.1252
21.7657 | -103.1271
21.7694 | -103.1281
21.7731 | -103.1261
21.7768 | -103.1227
21.7806 | -103.1168
21.7843 | -103.1096
21.788 | -103.1042
21.7918 | -103.0985
21.7955 | -103.0944
21.7992 | -103.0912
21.8029 | -103.0878
21.8067 | -103.0838
21.8104 | -103.0768
21.8141 | -103.0721
21.8178 | -103.0675
21.8216 | -103.0636
21.8253 | -103.0606
21.829 | -103.0575
21.8328 | -103.0545
21.8365 | -103.0515
21.8402 | -103.0482
21.8439 | -103.0437
21.8477 | -103.0383
21.8514 | -103.0349
21.8551 | -103.0302
21.8588 | -103.0246
21.8626 | -103.0203
21.8663 | -103.0178
21.87 | -103.0132
21.8737 | -103.0107
21.8775 | -103.0076
21.8812 | -103.0043
21.8849 | -103.0008
21.8887 | -102.9978
21.8924 | -102.9932
21.8961 | -102.9879
21.8998 | -102.9825
21.9036 | -102.9787
21.9073 | -102.977
21.911 | -102.9749
21.9147 | -102.9716
21.9185 | -102.969
21.9222 | -102.9661
21.9259 | -102.9635
21.9297 | -102.9625
21.9334 | -102.9623
21.9371 | -102.962
21.9408 | -102.9618
21.9446 | -102.9634
21.9483 | -102.9658
21.952 | -102.9683
21.9557 | -102.9716
21.9595 | -102.9765
21.9632 | -102.9822
21.9669 | -102.9874
21.9706 | -102.9937
21.9744 | -103.0008
21.9781 | -103.0083
21.9818 | -103.017
21.9856 | -103.0266
21.9893 | -103.0342
21.993 | -103.0396
21.9967 | -103.0463
22.0005 | -103.0542
22.0042 | -103.0633
22.0079 | -103.0734
22.0116 | -103.0824
22.0154 | -103.0922
22.0191 | -103.1016
22.0228 | -103.1112
22.0266 | -103.1209
22.0303 | -103.1306
22.034 | -103.1406
22.0377 | -103.151
22.0415 | -103.1608
22.0452 | -103.1669
22.0489 | -103.1757
22.0526 | -103.1854
22.0564 | -103.1954
22.0601 | -103.2028
22.0638 | -103.2112
22.0676 | -103.2207
22.0713 | -103.2311
22.075 | -103.2392
22.0787 | -103.2441
22.0825 | -103.25
22.0862 | -103.2591
22.0899 | -103.2678
22.0936 | -103.2772
22.0974 | -103.2876
22.1011 | -103.2964
22.1048 | -103.306
22.1085 | -103.311
22.1123 | -103.3154
22.116 | -103.3229
22.1197 | -103.3292
22.1235 | -103.3364
22.1272 | -103.3436
22.1309 | -103.3506
22.1346 | -103.3552
22.1384 | -103.3577
22.1421 | -103.3611
22.1458 | -103.366
22.1495 | -103.3692
22.1533 | -103.3723
22.157 | -103.3749
22.1607 | -103.3774
22.1645 | -103.3784
22.1682 | -103.3808
22.1719 | -103.3822
22.1756 | -103.3826
22.1794 | -103.3816
22.1831 | -103.3828
22.1868 | -103.3813
22.1905 | -103.3791
22.1943 | -103.3769
22.198 | -103.376
22.2017 | -103.3743
22.2055 | -103.3734
22.2092 | -103.372
22.2129 | -103.371
22.2166 | -103.3681
22.2204 | -103.3628
22.2241 | -103.3562
22.2278 | -103.3497
22.2315 | -103.3422
22.2353 | -103.3353
22.239 | -103.3304
22.2427 | -103.3263
22.2464 | -103.3198
22.2502 | -103.3129
22.2539 | -103.3067
22.2576 | -103.3002
22.2614 | -103.2933
22.2651 | -103.2873
22.2688 | -103.2816
22.2725 | -103.2785
22.2763 | -103.2749
22.28 | -103.2713
22.2837 | -103.2694
22.2874 | -103.2673
22.2912 | -103.2656
22.2949 | -103.2646
22.2986 | -103.2637
22.3024 | -103.2634
22.3061 | -103.2635
22.3098 | -103.2638
22.3135 | -103.265
22.3173 | -103.2661
22.321 | -103.2674
22.3247 | -103.2704
22.3284 | -103.2753
22.3322 | -103.2792
22.3359 | -103.2833
22.3396 | -103.2887
22.3434 | -103.2933
22.3471 | -103.2963
22.3508 | -103.3005
22.3545 | -103.3085
22.3583 | -103.3177
22.362 | -103.3251
22.3657 | -103.3326
22.3694 | -103.3423
22.3732 | -103.3509
22.3769 | -103.3593
22.3806 | -103.3662
22.3843 | -103.3731
22.3881 | -103.3814
22.3918 | -103.3904
22.3955 | -103.3989
22.3993 | -103.4071
22.403 | -103.4139
22.4067 | -103.4192
22.4104 | -103.4235
22.4142 | -103.4267
22.4179 | -103.43
22.4216 | -103.4304
22.4253 | -103.4309
22.4291 | -103.4296
22.4328 | -103.4266
22.4365 | -103.4245
22.4403 | -103.4232
22.444 | -103.419
22.4477 | -103.4133
22.4514 | -103.4072
22.4552 | -103.4011
22.4589 | -103.3927
22.4626 | -103.383
22.4663 | -103.3739
22.4701 | -103.3641
22.4738 | -103.3525
22.4775 | -103.3394
22.4812 | -103.325
22.485 | -103.3105
22.4887 | -103.2958
22.4924 | -103.2791
22.4962 | -103.2618
22.4999 | -103.243
22.5036 | -103.2268
22.5073 | -103.2086
22.5111 | -103.1892
22.5148 | -103.1697
22.5185 | -103.1503
22.5222 | -103.1313
22.526 | -103.1131
22.5297 | -103.0943
22.5334 | -103.0777
22.5372 | -103.0625
22.5409 | -103.0471
22.5446 | -103.0333
22.5483 | -103.0223
22.5521 | -103.012
22.5558 | -103.0004
22.5595 | -102.9906
22.5632 | -102.9836
22.567 | -102.9782
22.5707 | -102.9739
22.5744 | -102.9705
22.5782 | -102.9679
22.5819 | -102.9837
22.5856 | -102.9868
22.5893 | -102.9894
22.5931 | -102.9914
22.5968 | -102.9969
22.6005 | -103.0047
22.6042 | -103.0129
22.608 | -103.0239
22.6117 | -103.0386
22.6154 | -103.0533
22.6191 | -103.0689
22.6229 | -103.0884
22.6266 | -103.1098
22.6303 | -103.1307
22.6341 | -103.1525
22.6378 | -103.1751
22.6415 | -103.1979
22.6452 | -103.2195
22.649 | -103.2401
22.6527 | -103.2635
22.6564 | -103.2884
22.6601 | -103.312
22.6639 | -103.3374
22.6676 | -103.3635
22.6713 | -103.3894
22.6751 | -103.4131
22.6788 | -103.4364
22.6825 | -103.4609
22.6862 | -103.4847
22.69 | -103.5059
22.6937 | -103.5267
22.6974 | -103.547
22.7011 | -103.5672
22.7049 | -103.5874
22.7086 | -103.6072
22.7123 | -103.6249
22.716 | -103.644
22.7198 | -103.6615
22.7235 | -103.6776
22.7272 | -103.6951
22.731 | -103.7116
22.7347 | -103.7264
22.7384 | -103.7394
22.7421 | -103.7535
22.7459 | -103.7682
22.7496 | -103.7823
22.7533 | -103.7943
22.757 | -103.8047
22.7608 | -103.8126
22.7645 | -103.8213
22.7682 | -103.8272
22.772 | -103.8344
22.7757 | -103.8413
22.7794 | -103.8473
22.7831 | -103.8521
22.7869 | -103.857
22.7906 | -103.8606
22.7943 | -103.8648
22.798 | -103.8697
22.8018 | -103.874
22.8055 | -103.8768
22.8092 | -103.8775
22.813 | -103.8811
22.8167 | -103.8858
22.8204 | -103.888
22.8241 | -103.8877
22.8279 | -103.8896
22.8316 | -103.8914
22.8353 | -103.8916
22.839 | -103.8918
22.8428 | -103.892
22.8465 | -103.8908
22.8502 | -103.8899
22.8539 | -103.8879
22.8577 | -103.8875
22.8614 | -103.8875
22.8651 | -103.8866
22.8689 | -103.8831
22.8726 | -103.8773
22.8763 | -103.8725
22.88 | -103.8661
22.8838 | -103.8609
22.8875 | -103.8565
22.8912 | -103.853
22.8949 | -103.8496
22.8987 | -103.8471
22.9024 | -103.841
22.9061 | -103.8348
22.9099 | -103.8283
22.9136 | -103.8248
22.9173 | -103.8227
22.921 | -103.8189
22.9248 | -103.8141
22.9285 | -103.8093
22.9322 | -103.8024
22.9359 | -103.7937
22.9397 | -103.7879
22.9434 | -103.7836
22.9471 | -103.7793
22.9508 | -103.7759
22.9546 | -103.7737
22.9583 | -103.7705
22.962 | -103.7686
22.9658 | -103.7683
22.9695 | -103.7693
22.9732 | -103.7698
22.9769 | -103.7699
22.9807 | -103.7693
22.9844 | -103.7699
22.9881 | -103.7679
22.9918 | -103.7695
22.9956 | -103.7742
22.9993 | -103.7786
23.003 | -103.7834
23.0068 | -103.7892
23.0105 | -103.7943
23.0142 | -103.8007
23.0179 | -103.8075
23.0217 | -103.8157
23.0254 | -103.8246
23.0291 | -103.8349
23.0328 | -103.8437
23.0366 | -103.8534
23.0403 | -103.8635
23.044 | -103.8743
23.0478 | -103.8854
23.0515 | -103.8955
23.0552 | -103.9057
23.0589 | -103.9167
23.0627 | -103.9294
23.0664 | -103.94
23.0701 | -103.9511
23.0738 | -103.9612
23.0776 | -103.9698
23.0813 | -103.9802
23.085 | -103.9896
23.0887 | -103.9977
23.0925 | -104.007
23.0962 | -104.0147
23.0999 | -104.0201
23.1037 | -104.0254
23.1074 | -104.0323
23.1111 | -104.038
23.1148 | -104.0434
23.1186 | -104.0443
23.1223 | -104.0476
23.126 | -104.0484
23.1297 | -104.0512
23.1335 | -104.0546
23.1372 | -104.057
23.1409 | -104.0587
23.1447 | -104.0584
23.1484 | -104.0589
23.1521 | -104.0592
23.1558 | -104.058
23.1596 | -104.0572
23.1633 | -104.0572
23.167 | -104.0555
23.1707 | -104.0519
23.1745 | -104.0495
23.1782 | -104.0471
23.1819 | -104.0454
23.1857 | -104.043
23.1894 | -104.042
23.1931 | -104.0417
23.1968 | -104.0427
23.2006 | -104.0443
23.2043 | -104.0456
23.208 | -104.0482
23.2117 | -104.0487
23.2155 | -104.048
23.2192 | -104.0477
23.2229 | -104.0491
23.2266 | -104.052
23.2304 | -104.0565
23.2341 | -104.0601
23.2378 | -104.0638
23.2416 | -104.0667
23.2453 | -104.0691
23.249 | -104.0719
23.2527 | -104.0754
23.2565 | -104.0793
23.2602 | -104.0844
23.2639 | -104.0891
23.2676 | -104.0936
23.2714 | -104.0967
23.2751 | -104.0995
23.2788 | -104.1051
23.2826 | -104.1094
23.2863 | -104.1149
23.29 | -104.1199
23.2937 | -104.1251
23.2975 | -104.1293
23.3012 | -104.1341
23.3049 | -104.14
23.3086 | -104.1455
23.3124 | -104.1493
23.3161 | -104.1561
23.3198 | -104.1614
23.3236 | -104.1644
23.3273 | -104.1682
23.331 | -104.1731
23.3347 | -104.1787
23.3385 | -104.1839
23.3422 | -104.1869
23.3459 | -104.1903
23.3496 | -104.1951
23.3534 | -104.1987
23.3571 | -104.1984
23.3608 | -104.2
23.3645 | -104.2024
23.3683 | -104.2044
23.372 | -104.2055
23.3757 | -104.2074
23.3795 | -104.2102
23.3832 | -104.2131
23.3869 | -104.2145
23.3906 | -104.2149
23.3944 | -104.2171
23.3981 | -104.2184
23.4018 | -104.2199
23.4055 | -104.2221
23.4093 | -104.223
23.413 | -104.2235
23.4167 | -104.2248
23.4205 | -104.2268
23.4242 | -104.2288
23.4279 | -104.2311
23.4316 | -104.2332
23.4354 | -104.2344
23.4391 | -104.2362
23.4428 | -104.2368
23.4465 | -104.2366
23.4503 | -104.2361
23.454 | -104.2368
23.4577 | -104.2388
23.4614 | -104.2393
23.4652 | -104.2397
23.4689 | -104.241
23.4726 | -104.241
23.4764 | -104.2428
23.4801 | -104.2457
23.4838 | -104.2502
23.4875 | -104.2543
23.4913 | -104.2574
23.495 | -104.2607
23.4987 | -104.2661
23.5024 | -104.2729
23.5062 | -104.2795
23.5099 | -104.2853
23.5136 | -104.289
23.5174 | -104.2952
23.5211 | -104.3011
23.5248 | -104.3079
23.5285 | -104.3143
23.5323 | -104.3231
23.536 | -104.3335
23.5397 | -104.343
23.5434 | -104.3508
23.5472 | -104.3599
23.5509 | -104.3706
23.5546 | -104.3806
23.5584 | -104.3902
23.5621 | -104.4014
23.5658 | -104.4132
23.5695 | -104.427
23.5733 | -104.4424
23.577 | -104.4582
23.5807 | -104.4729
23.5844 | -104.4872
23.5882 | -104.501
23.5919 | -104.5154
23.5956 | -104.5308
23.5993 | -104.5452
23.6031 | -104.5597
23.6068 | -104.5748
23.6105 | -104.5878
23.6143 | -104.6016
23.618 | -104.6141
23.6217 | -104.6262
23.6254 | -104.6393
23.6292 | -104.6536
23.6329 | -104.666
23.6366 | -104.6762
23.6403 | -104.6873
23.6441 | -104.6985
23.6478 | -104.7073
23.6515 | -104.7161
23.6553 | -104.7223
23.659 | -104.7286
23.6627 | -104.7352
23.6664 | -104.7415
23.6702 | -104.7459
23.6739 | -104.7514
23.6776 | -104.759
23.6813 | -104.766
23.6851 | -104.7717
23.6888 | -104.7786
23.6925 | -104.787
23.6962 | -104.7946
23.7 | -104.7993
23.7037 | -104.8036
23.7074 | -104.8086
23.7112 | -104.8145
23.7149 | -104.8192
23.7186 | -104.8238
23.7223 | -104.8278
23.7261 | -104.8309
23.7298 | -104.833
23.7335 | -104.8338
23.7372 | -104.8334
23.741 | -104.832
23.7447 | -104.8319
23.7484 | -104.8352
23.7522 | -104.8367
23.7559 | -104.8363
23.7596 | -104.8357
23.7633 | -104.8353
23.7671 | -104.8339
23.7708 | -104.8326
23.7745 | -104.8309
23.7782 | -104.8285
23.782 | -104.8253
23.7857 | -104.8244
23.7894 | -104.8232
23.7932 | -104.8201
23.7969 | -104.8151
23.8006 | -104.8115
23.8043 | -104.807
23.8081 | -104.8033
23.8118 | -104.7981
23.8155 | -104.7916
23.8192 | -104.7843
23.823 | -104.7775
23.8267 | -104.7731
23.8304 | -104.7661
23.8341 | -104.7595
23.8379 | -104.7532
23.8416 | -104.7468
23.8453 | -104.7417
23.8491 | -104.7372
23.8528 | -104.7281
23.8565 | -104.7184
23.8602 | -104.7118
23.864 | -104.7067
23.8677 | -104.6999
23.8714 | -104.6932
23.8751 | -104.6875
23.8789 | -104.6856
23.8826 | -104.6832
23.8863 | -104.6768
23.8901 | -104.6695
23.8938 | -104.6637
23.8975 | -104.6571
23.9012 | -104.6509
23.905 | -104.645
23.9087 | -104.6384
23.9124 | -104.6317
23.9161 | -104.6275
23.9199 | -104.6241
23.9236 | -104.62
23.9273 | -104.6161
23.9311 | -104.6098
23.9348 | -104.6038
23.9385 | -104.599
23.9422 | -104.5954
23.946 | -104.5909
23.9497 | -104.5858
23.9534 | -104.5832
23.9571 | -104.5812
23.9609 | -104.5768
23.9646 | -104.5727
23.9683 | -104.5702
23.972 | -104.5685
23.9758 | -104.5665
23.9795 | -104.5652
23.9832 | -104.5639
23.987 | -104.5612
23.9907 | -104.558
23.9944 | -104.5569
23.9981 | -104.5531
24.0019 | -104.5477
24.0056 | -104.5445
24.0093 | -104.543
24.013 | -104.5413
24.0168 | -104.5379
24.0205 | -104.5353
24.0242 | -104.5327
24.028 | -104.5288
24.0317 | -104.5247
24.0354 | -104.5197
24.0391 | -104.5167
24.0429 | -104.5144
24.0466 | -104.5124
24.0503 | -104.5073
24.054 | -104.5029
24.0578 | -104.4994
24.0615 | -104.4946
24.0652 | -104.4916
24.0689 | -104.4905
24.0727 | -104.4892
24.0764 | -104.4896
24.0801 | -104.4909
24.0839 | -104.489
24.0876 | -104.4851
24.0913 | -104.4822
24.095 | -104.4818
24.0988 | -104.4793
24.1025 | -104.4774
24.1062 | -104.4761
24.1099 | -104.4748
24.1137 | -104.4732
24.1174 | -104.4728
24.1211 | -104.4737
24.1249 | -104.473
24.1286 | -104.4721
24.1323 | -104.4712
24.136 | -104.4709
24.1398 | -104.4697
24.1435 | -104.4709
24.1472 | -104.469
24.1509 | -104.4679
24.1547 | -104.4697
24.1584 | -104.4708
24.1621 | -104.4721
24.1659 | -104.4702
24.1696 | -104.4691
24.1733 | -104.4691
24.177 | -104.4684
24.1808 | -104.4668
24.1845 | -104.4666
24.1882 | -104.4682
24.1919 | -104.4722
24.1957 | -104.4762
24.1994 | -104.4804
24.2031 | -104.4814
24.2068 | -104.4829
24.2106 | -104.486
24.2143 | -104.4888
24.218 | -104.4921
24.2218 | -104.4959
24.2255 | -104.5006
24.2292 | -104.5079
24.2329 | -104.5154
24.2367 | -104.52
24.2404 | -104.5252
24.2441 | -104.5281
24.2478 | -104.5256
24.2516 | -104.519
24.2553 | -104.509
24.259 | -104.498
24.2628 | -104.4723
24.2665 | -104.4346
24.2702 | -104.3953
24.2739 | -104.3646
24.2777 | -104.3379
24.2814 | -104.3103
24.2851 | -104.2832
24.2888 | -104.2561
24.2926 | -104.2284
24.2963 | -104.1965
24.3 | -104.1609
24.3038 | -104.1238
24.3075 | -104.0844
24.3112 | -104.0413
24.3149 | -103.9959
24.3187 | -103.9484
24.3224 | -103.8963
24.3261 | -103.8417
24.3298 | -103.7861
24.3336 | -103.7285
24.3373 | -103.6697
24.341 | -103.6123
24.3447 | -103.553
24.3485 | -103.4905
24.3522 | -103.4261
24.3559 | -103.3617
24.3597 | -103.2981
24.3634 | -103.2345
24.3671 | -103.1698
24.3708 | -103.1063
24.3746 | -103.0447
24.3783 | -102.9825
24.382 | -102.9219
24.3857 | -102.8638
24.3895 | -102.8068
24.3932 | -102.7545
24.3969 | -102.7046
24.4007 | -102.6579
24.4044 | -102.6142
24.4081 | -102.5728
24.4118 | -102.5356
24.4156 | -102.5014
24.4193 | -102.4706
24.423 | -102.444
24.4267 | -102.4217
24.4305 | -102.4036
24.4342 | -102.3894
24.4379 | -102.3791
24.4416 | -102.3728
24.4454 | -102.3713
24.4491 | -102.3726
24.4528 | -102.3769
24.4566 | -102.3851
24.4603 | -102.3976
24.464 | -102.4136
24.4677 | -102.4323
24.4715 | -102.4546
24.4752 | -102.4804
24.4789 | -102.509
24.4826 | -102.5406
24.4864 | -102.5732
24.4901 | -102.6095
24.4938 | -102.6467
24.4976 | -102.6839
24.5013 | -102.7228
24.505 | -102.7635
24.5087 | -102.8073
24.5125 | -102.8522
24.5162 | -102.8957
24.5199 | -102.9392
24.5236 | -102.9822
24.5274 | -103.0255
24.5311 | -103.069
24.5348 | -103.1118
24.5386 | -103.1539
24.5423 | -103.1963
24.546 | -103.2373
24.5497 | -103.2769
24.5535 | -103.3141
24.5572 | -103.3444
24.5609 | -103.3644
24.5646 | -103.3747
24.5684 | -103.3853
24.5721 | -103.3978
24.5758 | -103.4068
24.5795 | -103.4128
24.5833 | -103.4075
24.587 | -103.3636
24.5907 | -103.28
24.5945 | -103.1455
24.5982 | -103.0102
24.6019 | -102.9003
24.6056 | -102.8085
24.6094 | -102.7292
24.6131 | -102.6544
24.6168 | -102.5679
24.6205 | -102.4647
24.6243 | -102.3506
24.628 | -102.2296
24.6317 | -102.1013
24.6355 | -101.9681
24.6392 | -101.8316
24.6429 | -101.6915
24.6466 | -101.5501
24.6504 | -101.4066
24.6541 | -101.2608
24.6578 | -101.113
24.6615 | -100.9648
24.6653 | -100.8165
24.669 | -100.6703
24.6727 | -100.5247
24.6764 | -100.3801
24.6802 | -100.2377
24.6839 | -100.0975
24.6876 | -99.9605
24.6914 | -99.8271
24.6951 | -99.6971
24.6988 | -99.5724
24.7025 | -99.4528
24.7063 | -99.3379
24.71 | -99.2284
24.7137 | -99.1255
24.7174 | -99.0276
24.7212 | -98.9363
24.7249 | -98.8525
24.7286 | -98.7753
24.7324 | -98.7541
24.7361 | -98.6934
24.7398 | -98.64
24.7435 | -98.5932
24.7473 | -98.553
24.751 | -98.5199
24.7547 | -98.495
24.7584 | -98.4772
24.7622 | -98.4658
24.7659 | -98.4626
24.7696 | -98.4673
24.7734 | -98.4785
24.7771 | -98.4965
24.7808 | -98.5218
24.7845 | -98.5548
24.7883 | -98.5945
24.792 | -98.6422
24.7957 | -98.6968
24.7994 | -98.7577
24.8032 | -98.8253
24.8069 | -98.8998
24.8106 | -98.9803
24.8143 | -99.0666
24.8181 | -99.1594
24.8218 | -99.2578
24.8255 | -99.3613
24.8293 | -99.47
24.833 | -99.5834
24.8367 | -99.7006
24.8404 | -99.8217
24.8442 | -99.947
24.8479 | -100.0758
24.8516 | -100.2072
24.8553 | -100.3409
24.8591 | -100.4764
24.8628 | -100.6135
24.8665 | -100.7513
24.8703 | -100.89
24.874 | -101.0287
24.8777 | -101.1659
24.8814 | -101.3016
24.8852 | -101.4352
24.8889 | -101.5673
24.8926 | -101.6966
24.8963 | -101.8246
24.9001 | -101.9496
24.9038 | -102.0708
24.9075 | -102.1912
24.9113 | -102.3121
24.915 | -102.4316
24.9187 | -102.542
24.9224 | -102.6423
24.9262 | -102.7368
24.9299 | -102.8259
24.9336 | -102.9156
24.9373 | -103.025
24.9411 | -103.1509
24.9448 | -103.3045
24.9485 | -103.4352
24.9522 | -103.5158
24.956 | -103.5551
24.9597 | -103.575
24.9634 | -103.5841
24.9672 | -103.5908
24.9709 | -103.5957
24.9746 | -103.5982
24.9783 | -103.6005
24.9821 | -103.5997
24.9858 | -103.5967
24.9895 | -103.5936
24.9932 | -103.591
24.997 | -103.5858
25.0007 | -103.5787
25.0044 | -103.572
25.0082 | -103.569
25.0119 | -103.5655
25.0156 | -103.5616
25.0193 | -103.5569
25.0231 | -103.5516
25.0268 | -103.5467
25.0305 | -103.5427
25.0342 | -103.538
25.038 | -103.5313
25.0417 | -103.5262
25.0454 | -103.5218
25.0492 | -103.5165
25.0529 | -103.5099
25.0566 | -103.5032
25.0603 | -103.4981
25.0641 | -103.4932
25.0678 | -103.4926
25.0715 | -103.4905
25.0752 | -103.4877
25.079 | -103.4842
25.0827 | -103.4808
25.0864 | -103.4797
25.0901 | -103.4793
25.0939 | -103.4784
25.0976 | -103.4763
25.1013 | -103.4769
25.1051 | -103.4786
25.1088 | -103.48
25.1125 | -103.4822
25.1162 | -103.4854
25.12 | -103.489
25.1237 | -103.491
25.1274 | -103.4946
25.1311 | -103.5015
25.1349 | -103.5067
25.1386 | -103.5104
25.1423 | -103.5145
25.1461 | -103.5176
25.1498 | -103.5198
25.1535 | -103.5238
25.1572 | -103.5293
25.161 | -103.5336
25.1647 | -103.5391
25.1684 | -103.5437
25.1721 | -103.5471
25.1759 | -103.5517
25.1796 | -103.5579
25.1833 | -103.5628
25.187 | -103.5664
25.1908 | -103.5721
25.1945 | -103.5771
25.1982 | -103.5816
25.202 | -103.5876
25.2057 | -103.594
25.2094 | -103.5992
25.2131 | -103.6024
25.2169 | -103.6047
25.2206 | -103.6084
25.2243 | -103.6119
25.228 | -103.6155
25.2318 | -103.6186
25.2355 | -103.6216
25.2392 | -103.6258
25.243 | -103.63
25.2467 | -103.6336
25.2504 | -103.6347
25.2541 | -103.6362
25.2579 | -103.6375
25.2616 | -103.6388
25.2653 | -103.6419
25.269 | -103.6477
25.2728 | -103.6506
25.2765 | -103.6539
25.2802 | -103.6578
25.284 | -103.6611
25.2877 | -103.665
25.2914 | -103.6692
25.2951 | -103.673
25.2989 | -103.6815
25.3026 | -103.6885
25.3063 | -103.6923
25.31 | -103.6963
25.3138 | -103.7013
25.3175 | -103.7059
25.3212 | -103.7098
25.3249 | -103.7153
25.3287 | -103.7215
25.3324 | -103.7296
25.3361 | -103.7368
25.3399 | -103.7459
25.3436 | -103.7558
25.3473 | -103.7649
25.351 | -103.7742
25.3548 | -103.7848
25.3585 | -103.7958
25.3622 | -103.8075
25.3659 | -103.8175
25.3697 | -103.8257
25.3734 | -103.8355
25.3771 | -103.8457
25.3809 | -103.8569
25.3846 | -103.8686
25.3883 | -103.8794
25.392 | -103.8911
25.3958 | -103.9021
25.3995 | -103.9128
25.4032 | -103.924
25.4069 | -103.935
25.4107 | -103.9452
25.4144 | -103.9565
25.4181 | -103.9662
25.4218 | -103.975
25.4256 | -103.9825
25.4293 | -103.9906
25.433 | -103.998
25.4368 | -104.0077
25.4405 | -104.0163
25.4442 | -104.0234
25.4479 | -104.029
25.4517 | -104.0346
25.4554 | -104.04
25.4591 | -104.043
25.4628 | -104.0471
25.4666 | -104.0494
25.4703 | -104.0511
25.474 | -104.0527
25.4778 | -104.0535
25.4815 | -104.0541
25.4852 | -104.0541
25.4889 | -104.0531
25.4927 | -104.0531
25.4964 | -104.0523
25.5001 | -104.0502
25.5038 | -104.0476
25.5076 | -104.0442
25.5113 | -104.0413
25.515 | -104.0363
25.5188 | -104.0298
25.5225 | -104.0214
25.5262 | -104.0154
25.5299 | -104.0072
25.5337 | -103.9971
25.5374 | -103.9879
25.5411 | -103.9786
25.5448 | -103.9689
25.5486 | -103.9591
25.5523 | -103.9481
25.556 | -103.9374
25.5597 | -103.9255
25.5635 | -103.9138
25.5672 | -103.9038
25.5709 | -103.8936
25.5747 | -103.8836
25.5784 | -103.8734
25.5821 | -103.8635
25.5858 | -103.8543
25.5896 | -103.8444
25.5933 | -103.8344
25.597 | -103.8243
25.6007 | -103.815
25.6045 | -103.8047
25.6082 | -103.7961
25.6119 | -103.7881
25.6157 | -103.7819
25.6194 | -103.7765
25.6231 | -103.7692
25.6268 | -103.7613
25.6306 | -103.7544
25.6343 | -103.7514
25.638 | -103.7479
25.6417 | -103.7446
25.6455 | -103.7434
25.6492 | -103.7406
25.6529 | -103.7372
25.6566 | -103.7354
25.6604 | -103.7335
25.6641 | -103.732
25.6678 | -103.7325
25.6716 | -103.7336
25.6753 | -103.7347
25.679 | -103.7366
25.6827 | -103.74
25.6865 | -103.7428
25.6902 | -103.7474
25.6939 | -103.7502
25.6976 | -103.7534
25.7014 | -103.757
25.7051 | -103.7616
25.7088 | -103.7665
25.7126 | -103.772
25.7163 | -103.7773
25.72 | -103.7847
25.7237 | -103.791
25.7275 | -103.7962
25.7312 | -103.8044
25.7349 | -103.8116
25.7386 | -103.8189
25.7424 | -103.8269
25.7461 | -103.834
25.7498 | -103.8409
25.7536 | -103.85
25.7573 | -103.861
25.761 | -103.8689
25.7647 | -103.8764
25.7685 | -103.8845
25.7722 | -103.8935
25.7759 | -103.9011
25.7796 | -103.9089
25.7834 | -103.9171
25.7871 | -103.9255
25.7908 | -103.934
25.7945 | -103.9438
25.7983 | -103.9516
25.802 | -103.9584
25.8057 | -103.9635
25.8095 | -103.9716
25.8132 | -103.9794
25.8169 | -103.9856
25.8206 | -103.9933
25.8244 | -104.0013
25.8281 | -104.0062
25.8318 | -104.0128
25.8355 | -104.02
25.8393 | -104.0262
25.843 | -104.0299
25.8467 | -104.0348
25.8505 | -104.0415
25.8542 | -104.0476
25.8579 | -104.0528
25.8616 | -104.058
25.8654 | -104.0635
25.8691 | -104.0693
25.8728 | -104.0755
25.8765 | -104.0788
25.8803 | -104.0845
25.884 | -104.0898
25.8877 | -104.0946
25.8915 | -104.0988
25.8952 | -104.1062
25.8989 | -104.1117
25.9026 | -104.1146
25.9064 | -104.1188
25.9101 | -104.1241
25.9138 | -104.1289
25.9175 | -104.1345
25.9213 | -104.1401
25.925 | -104.1461
25.9287 | -104.1522
25.9324 | -104.1559
25.9362 | -104.1616
25.9399 | -104.1676
25.9436 | -104.1736
25.9474 | -104.1786
25.9511 | -104.1838
25.9548 | -104.1908
25.9585 | -104.1972
25.9623 | -104.2035
25.966 | -104.2088
25.9697 | -104.2169
25.9734 | -104.2238
25.9772 | -104.2299
25.9809 | -104.2367
25.9846 | -104.2433
25.9884 | -104.2517
25.9921 | -104.2597
25.9958 | -104.2673
25.9995 | -104.274
26.0033 | -104.2815
26.007 | -104.2887
26.0107 | -104.2949
26.0144 | -104.3024
26.0182 | -104.3115
26.0219 | -104.3223
26.0256 | -104.3331
26.0294 | -104.3449
26.0331 | -104.3543
26.0368 | -104.3643
26.0405 | -104.3751
26.0443 | -104.3856
26.048 | -104.3957
26.0517 | -104.4048
26.0554 | -104.4134
26.0592 | -104.4214
26.0629 | -104.4302
26.0666 | -104.4403
26.0703 | -104.4498
26.0741 | -104.458
26.0778 | -104.467
26.0815 | -104.4779
26.0853 | -104.4863
26.089 | -104.496
26.0927 | -104.5061
26.0964 | -104.5126
26.1002 | -104.5191
26.1039 | -104.5257
26.1076 | -104.5307
26.1113 | -104.5341
26.1151 | -104.5389
26.1188 | -104.5435
26.1225 | -104.547
26.1263 | -104.5522
26.13 | -104.5553
26.1337 | -104.5567
26.1374 | -104.5578
26.1412 | -104.5581
26.1449 | -104.5553
26.1486 | -104.5524
26.1523 | -104.547
26.1561 | -104.5417
26.1598 | -104.535
26.1635 | -104.5283
26.1672 | -104.5223
26.171 | -104.5163
26.1747 | -104.5084
26.1784 | -104.4997
26.1822 | -104.4903
26.1859 | -104.4818
26.1896 | -104.4717
26.1933 | -104.4602
26.1971 | -104.4486
26.2008 | -104.4349
26.2045 | -104.4203
26.2082 | -104.4069
26.212 | -104.3947
26.2157 | -104.3818
26.2194 | -104.367
26.2232 | -104.352
26.2269 | -104.337
26.2306 | -104.3221
26.2343 | -104.307
26.2381 | -104.2913
26.2418 | -104.2764
26.2455 | -104.2639
26.2492 | -104.2502
26.253 | -104.2383
26.2567 | -104.2261
26.2604 | -104.2146
26.2642 | -104.202
26.2679 | -104.1906
26.2716 | -104.1809
26.2753 | -104.1722
26.2791 | -104.164
26.2828 | -104.1555
26.2865 | -104.1494
26.2902 | -104.1432
26.294 | -104.1372
26.2977 | -104.1316
26.3014 | -104.1276
26.3051 | -104.1246
26.3089 | -104.1233
26.3126 | -104.1227
26.3163 | -104.1231
26.3201 | -104.1231
26.3238 | -104.1228
26.3275 | -104.1254
26.3312 | -104.1276
26.335 | -104.1319
26.3387 | -104.1365
26.3424 | -104.1416
26.3461 | -104.1529
26.3499 | -104.1584
26.3536 | -104.1634
26.3573 | -104.1689
26.3611 | -104.175
26.3648 | -104.1818
26.3685 | -104.1894
26.3722 | -104.197
26.376 | -104.2049
26.3797 | -104.2132
26.3834 | -104.2215
26.3871 | -104.2307
26.3909 | -104.2394
26.3946 | -104.2466
26.3983 | -104.2533
26.402 | -104.2602
26.4058 | -104.2689
26.4095 | -104.2774
26.4132 | -104.2833
26.417 | -104.2898
26.4207 | -104.2968
26.4244 | -104.3046
26.4281 | -104.3121
26.4319 | -104.321
26.4356 | -104.3279
26.4393 | -104.3333
26.443 | -104.3401
26.4468 | -104.3477
26.4505 | -104.3526
26.4542 | -104.3566
26.458 | -104.3611
26.4617 | -104.3655
26.4654 | -104.3706
26.4691 | -104.377
26.4729 | -104.3831
26.4766 | -104.3881
26.4803 | -104.3934
26.484 | -104.4003
26.4878 | -104.4052
26.4915 | -104.4075
26.4952 | -104.409
26.499 | -104.4126
26.5027 | -104.4181
26.5064 | -104.4215
26.5101 | -104.424
26.5139 | -104.4276
26.5176 | -104.4327
26.5213 | -104.4355
26.525 | -104.4376
26.5288 | -104.4396
26.5325 | -104.441
26.5362 | -104.441
26.5399 | -104.4415
26.5437 | -104.4425
26.5474 | -104.4425
26.5511 | -104.4406
26.5549 | -104.4406
26.5586 | -104.4406
26.5623 | -104.4425
26.566 | -104.4449
26.5698 | -104.4462
26.5735 | -104.4473
26.5772 | -104.4484
26.5809 | -104.4489
26.5847 | -104.4498
26.5884 | -104.4505
26.5921 | -104.4505
26.5959 | -104.4503
26.5996 | -104.4494
26.6033 | -104.4479
26.607 | -104.4479
26.6108 | -104.4479
26.6145 | -104.4456
26.6182 | -104.4433
26.6219 | -104.4421
26.6257 | -104.4407
26.6294 | -104.4404
26.6331 | -104.4376
26.6369 | -104.4351
26.6406 | -104.434
26.6443 | -104.4329
26.648 | -104.4317
26.6518 | -104.4299
26.6555 | -104.429
26.6592 | -104.4263
26.6629 | -104.4246
26.6667 | -104.4231
26.6704 | -104.4218
26.6741 | -104.4209
26.6778 | -104.4172
26.6816 | -104.4145
26.6853 | -104.4124
26.689 | -104.4101
26.6928 | -104.409
26.6965 | -104.4078
26.7002 | -104.4065
26.7039 | -104.4062
26.7077 | -104.405
26.7114 | -104.4044
26.7151 | -104.4041
26.7188 | -104.4013
26.7226 | -104.4004
26.7263 | -104.4022
26.73 | -104.403
26.7338 | -104.4028
26.7375 | -104.4025
26.7412 | -104.4036
26.7449 | -104.4027
26.7487 | -104.4016
26.7524 | -104.4017
26.7561 | -104.4014
26.7598 | -104.4007
26.7636 | -104.4015
26.7673 | -104.4038
26.771 | -104.4057
26.7748 | -104.4052
26.7785 | -104.4062
26.7822 | -104.4086
26.7859 | -104.4094
26.7897 | -104.412
26.7934 | -104.4149
26.7971 | -104.4174
26.8008 | -104.4208
26.8046 | -104.4226
26.8083 | -104.4237
26.812 | -104.424
26.8157 | -104.4252
26.8195 | -104.4267
26.8232 | -104.4278
26.8269 | -104.4291
26.8307 | -104.4298
26.8344 | -104.4314
26.8381 | -104.4326
26.8418 | -104.4328
26.8456 | -104.4326
26.8493 | -104.4325
26.853 | -104.4311
26.8567 | -104.4299
26.8605 | -104.4287
26.8642 | -104.428
26.8679 | -104.4277
26.8717 | -104.4274
26.8754 | -104.4247
26.8791 | -104.4215
26.8828 | -104.4173
26.8866 | -104.4155
26.8903 | -104.4118
26.894 | -104.4077
26.8977 | -104.404
26.9015 | -104.4002
26.9052 | -104.3969
26.9089 | -104.3909
26.9126 | -104.3847
26.9164 | -104.3803
26.9201 | -104.3742
26.9238 | -104.3655
26.9276 | -104.3589
26.9313 | -104.3547
26.935 | -104.3495
26.9387 | -104.3437
26.9425 | -104.337
26.9462 | -104.3293
26.9499 | -104.3216
26.9536 | -104.3154
26.9574 | -104.3081
26.9611 | -104.2997
26.9648 | -104.2918
26.9686 | -104.2833
26.9723 | -104.2764
26.976 | -104.2701
26.9797 | -104.2647
26.9835 | -104.2607
26.9872 | -104.2555
26.9909 | -104.2492
26.9946 | -104.2447
26.9984 | -104.24
27.0021 | -104.2346
27.0058 | -104.2291
27.0096 | -104.2228
27.0133 | -104.2164
27.017 | -104.212
27.0207 | -104.2081
27.0245 | -104.2029
27.0282 | -104.1967
27.0319 | -104.1912
27.0356 | -104.1862
27.0394 | -104.1813
27.0431 | -104.1762
27.0468 | -104.1718
27.0505 | -104.1669
27.0543 | -104.1624
27.058 | -104.1605
27.0617 | -104.157
27.0655 | -104.1543
27.0692 | -104.1519
27.0729 | -104.1486
27.0766 | -104.1436
27.0804 | -104.1383
27.0841 | -104.1335
27.0878 | -104.1267
27.0915 | -104.1203
27.0953 | -104.1154
27.099 | -104.1119
27.1027 | -104.1069
27.1065 | -104.1012
27.1102 | -104.0976
27.1139 | -104.0944
27.1176 | -104.0907
27.1214 | -104.0857
27.1251 | -104.0829
27.1288 | -104.0785
27.1325 | -104.0719
27.1363 | -104.0669
27.14 | -104.0629
27.1437 | -104.0568
27.1474 | -104.0498
27.1512 | -104.0443
27.1549 | -104.0419
27.1586 | -104.0382
27.1624 | -104.0334
27.1661 | -104.0293
27.1698 | -104.0245
27.1735 | -104.0191
27.1773 | -104.0142
27.181 | -104.0106
27.1847 | -104.0084
27.1884 | -104.0043
27.1922 | -104.0017
27.1959 | -103.9972
27.1996 | -103.9919
27.2034 | -103.9859
27.2071 | -103.9811
27.2108 | -103.977
27.2145 | -103.9732
27.2183 | -103.9692
27.222 | -103.965
27.2257 | -103.9607
27.2294 | -103.9568
27.2332 | -103.9534
27.2369 | -103.9489
27.2406 | -103.9425
27.2444 | -103.9376
27.2481 | -103.9323
27.2518 | -103.9274
27.2555 | -103.9219
27.2593 | -103.9176
27.263 | -103.913
27.2667 | -103.909
27.2704 | -103.9053
27.2742 | -103.9019
27.2779 | -103.8976
27.2816 | -103.8937
27.2853 | -103.8881
27.2891 | -103.8831
27.2928 | -103.8799
27.2965 | -103.8765
27.3003 | -103.8711
27.304 | -103.866
27.3077 | -103.863
27.3114 | -103.8564
27.3152 | -103.8479
27.3189 | -103.8393
27.3226 | -103.8311
27.3263 | -103.8236
27.3301 | -103.817
27.3338 | -103.8111
27.3375 | -103.8041
27.3413 | -103.7969
27.345 | -103.7915
27.3487 | -103.7828
27.3524 | -103.7727
27.3562 | -103.7643
27.3599 | -103.7577
27.3636 | -103.7505
27.3673 | -103.7421
27.3711 | -103.735
27.3748 | -103.7276
27.3785 | -103.721
27.3822 | -103.7146
27.386 | -103.7095
27.3897 | -103.7055
27.3934 | -103.7016
27.3972 | -103.6983
27.4009 | -103.6944
27.4046 | -103.6908
27.4083 | -103.688
27.4121 | -103.6847
27.4158 | -103.6821
27.4195 | -103.6863
27.4232 | -103.6858
27.427 | -103.6864
27.4307 | -103.6866
27.4344 | -103.6879
27.4382 | -103.6909
27.4419 | -103.6932
27.4456 | -103.696
27.4493 | -103.7003
27.4531 | -103.7038
27.4568 | -103.7077
27.4605 | -103.7135
27.4642 | -103.7195
27.468 | -103.7242
27.4717 | -103.7297
27.4754 | -103.737
27.4792 | -103.7442
27.4829 | -103.7491
27.4866 | -103.7555
27.4903 | -103.762
27.4941 | -103.7671
27.4978 | -103.7717
27.5015 | -103.7756
27.5052 | -103.7813
27.509 | -103.787
27.5127 | -103.7923
27.5164 | -103.7999
27.5201 | -103.8066
27.5239 | -103.8117
27.5276 | -103.8161
27.5313 | -103.8209
27.5351 | -103.8257
27.5388 | -103.8294
27.5425 | -103.8332
27.5462 | -103.8368
27.55 | -103.8397
27.5537 | -103.8432
27.5574 | -103.8469
27.5611 | -103.8493
27.5649 | -103.8496
27.5686 | -103.8493
27.5723 | -103.8487
27.5761 | -103.8502
27.5798 | -103.8502
27.5835 | -103.8496
27.5872 | -103.8493
27.591 | -103.8496
27.5947 | -103.8487
27.5984 | -103.8453
27.6021 | -103.8405
27.6059 | -103.8359
27.6096 | -103.83
27.6133 | -103.8258
27.6171 | -103.8211
27.6208 | -103.8165
27.6245 | -103.8133
27.6282 | -103.8102
27.632 | -103.8038
27.6357 | -103.7989
27.6394 | -103.7946
27.6431 | -103.7912
27.6469 | -103.7856
27.6506 | -103.7806
27.6543 | -103.776
27.658 | -103.77
27.6618 | -103.7648
27.6655 | -103.7603
27.6692 | -103.7551
27.673 | -103.7533
27.6767 | -103.7508
27.6804 | -103.7475
27.6841 | -103.7453
27.6879 | -103.7433
27.6916 | -103.7409
27.6953 | -103.739
27.699 | -103.7384
27.7028 | -103.7386
27.7065 | -103.7386
27.7102 | -103.7384
27.714 | -103.7372
27.7177 | -103.7361
27.7214 | -103.7362
27.7251 | -103.7365
27.7289 | -103.7367
27.7326 | -103.7373
27.7363 | -103.7412
27.74 | -103.7464
27.7438 | -103.7494
27.7475 | -103.7525
27.7512 | -103.7563
27.755 | -103.7581
27.7587 | -103.7593
27.7624 | -103.7623
27.7661 | -103.7668
27.7699 | -103.7699
27.7736 | -103.7721
27.7773 | -103.7743
27.781 | -103.7766
27.7848 | -103.7792
27.7885 | -103.7788
27.7922 | -103.7791
27.7959 | -103.7805
27.7997 | -103.7842
27.8034 | -103.7849
27.8071 | -103.7842
27.8109 | -103.783
27.8146 | -103.7798
27.8183 | -103.7763
27.822 | -103.7739
27.8258 | -103.7705
27.8295 | -103.7667
27.8332 | -103.7637
27.8369 | -103.7598
27.8407 | -103.7543
27.8444 | -103.7484
27.8481 | -103.7428
27.8519 | -103.735
27.8556 | -103.7279
27.8593 | -103.7223
27.863 | -103.7156
27.8668 | -103.7097
27.8705 | -103.7019
27.8742 | -103.6938
27.8779 | -103.6842
27.8817 | -103.6764
27.8854 | -103.6692
27.8891 | -103.6607
27.8928 | -103.6528
27.8966 | -103.6441
27.9003 | -103.6359
27.904 | -103.6279
27.9078 | -103.6195
27.9115 | -103.6134
27.9152 | -103.6068
27.9189 | -103.5988
27.9227 | -103.5903
27.9264 | -103.5799
27.9301 | -103.565
27.9338 | -103.5506
27.9376 | -103.5401
27.9413 | -103.5318
27.945 | -103.5237
27.9488 | -103.5165
27.9525 | -103.5108
27.9562 | -103.5063
27.9599 | -103.5006
27.9637 | -103.4953
27.9674 | -103.4909
27.9711 | -103.4868
27.9748 | -103.4834
27.9786 | -103.4807
27.9823 | -103.4765
27.986 | -103.4723
27.9898 | -103.4694
27.9935 | -103.4667
27.9972 | -103.4605
28.0009 | -103.4575