## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 271 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 263
Hood | 02-92B/P | 205
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1009.2 | hPa |  | PRC PRO 01
Humidity | 56.7 | %/Hr
Internal temperature of calibration bench | 24.21 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1017 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.6 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.04 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.01 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3122 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 23/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 21/08/2018
Technical Manager | Anthony HUE | Approval |  | 30/08/2018
CEO | Robert GLEMOT | Approval |  | 30/08/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3503 | 57.4095
0.012209 | -35.9457 | 38.2365
0.018313 | -34.8957 | 28.0741
0.024417 | -34.436 | 21.842
0.030521 | -34.3218 | 18.0637
0.036626 | -34.1793 | 15.0105
0.04273 | -34.0852 | 12.6575
0.048834 | -34.0403 | 11.1338
0.054938 | -33.9786 | 9.997
0.061043 | -33.9211 | 8.9261
0.067147 | -33.9116 | 8.11
0.073251 | -33.9123 | 7.3561
0.079355 | -33.9292 | 6.6994
0.08546 | -33.9424 | 6.1098
0.091564 | -33.9382 | 5.6692
0.097668 | -33.9216 | 5.4133
0.10377 | -33.9011 | 5.1647
0.10988 | -33.8852 | 4.8889
0.11598 | -33.8767 | 4.6001
0.12209 | -33.8784 | 4.3197
0.12819 | -33.8879 | 4.0394
0.13429 | -33.8983 | 3.7555
0.1404 | -33.9017 | 3.4847
0.1465 | -33.9038 | 3.2332
0.15261 | -33.9017 | 3.0178
0.15871 | -33.8962 | 2.8568
0.16482 | -33.8892 | 2.7511
0.17092 | -33.8874 | 2.7052
0.17702 | -33.8857 | 2.6288
0.18313 | -33.8851 | 2.5476
0.18923 | -33.8848 | 2.4382
0.19534 | -33.8837 | 2.3128
0.20144 | -33.8825 | 2.1736
0.20754 | -33.8801 | 2.0558
0.21365 | -33.8794 | 1.9436
0.21975 | -33.8792 | 1.8473
0.22586 | -33.8805 | 1.7611
0.23196 | -33.8831 | 1.6891
0.23807 | -33.8872 | 1.6261
0.24417 | -33.8913 | 1.5674
0.25027 | -33.8952 | 1.5052
0.25638 | -33.8977 | 1.4436
0.26248 | -33.8996 | 1.3777
0.26859 | -33.9007 | 1.3133
0.27469 | -33.9012 | 1.2435
0.2808 | -33.9023 | 1.1714
0.2869 | -33.9032 | 1.1018
0.293 | -33.9042 | 1.0358
0.29911 | -33.9051 | 0.97623
0.30521 | -33.9062 | 0.91764
0.31132 | -33.9069 | 0.87827
0.31742 | -33.9082 | 0.83509
0.32353 | -33.9086 | 0.79935
0.32963 | -33.9095 | 0.76927
0.33573 | -33.9106 | 0.74755
0.34184 | -33.912 | 0.72503
0.34794 | -33.9141 | 0.7008
0.35405 | -33.9163 | 0.67612
0.36015 | -33.9187 | 0.65042
0.36626 | -33.9207 | 0.62283
0.37236 | -33.9228 | 0.58672
0.37846 | -33.925 | 0.54367
0.38457 | -33.9259 | 0.48859
0.39067 | -33.9271 | 0.44325
0.39678 | -33.9278 | 0.39299
0.40288 | -33.9284 | 0.34066
0.40899 | -33.9284 | 0.29036
0.41509 | -33.9279 | 0.24098
0.42119 | -33.928 | 0.19871
0.4273 | -33.9281 | 0.15367
0.4334 | -33.9281 | 0.10876
0.43951 | -33.9285 | 0.066437
0.44561 | -33.9294 | 0.027507
0.45172 | -33.931 | -0.0037986
0.45782 | -33.9318 | -0.040163
0.46392 | -33.9331 | -0.075661
0.47003 | -33.9341 | -0.10758
0.47613 | -33.935 | -0.14113
0.48224 | -33.9356 | -0.17521
0.48834 | -33.9348 | -0.20919
0.49445 | -33.9347 | -0.24371
0.50055 | -33.9342 | -0.27522
0.50665 | -33.9336 | -0.30867
0.51276 | -33.9327 | -0.34193
0.51886 | -33.9318 | -0.3751
0.52497 | -33.9308 | -0.39903
0.53107 | -33.9301 | -0.42796
0.53717 | -33.9295 | -0.4556
0.54328 | -33.9288 | -0.4783
0.54938 | -33.9287 | -0.50063
0.55549 | -33.9288 | -0.52155
0.56159 | -33.9299 | -0.53377
0.5677 | -33.9303 | -0.54874
0.5738 | -33.9314 | -0.56533
0.5799 | -33.9322 | -0.58067
0.58601 | -33.9333 | -0.59154
0.59211 | -33.9349 | -0.599
0.59822 | -33.9362 | -0.60934
0.60432 | -33.9377 | -0.62215
0.61043 | -33.939 | -0.63108
0.61653 | -33.9402 | -0.64219
0.62263 | -33.9417 | -0.65241
0.62874 | -33.9423 | -0.66373
0.63484 | -33.9432 | -0.67381
0.64095 | -33.9437 | -0.6847
0.64705 | -33.9441 | -0.69518
0.65316 | -33.9444 | -0.70813
0.65926 | -33.9445 | -0.72419
0.66536 | -33.9435 | -0.74158
0.67147 | -33.9433 | -0.75508
0.67757 | -33.943 | -0.76963
0.68368 | -33.9423 | -0.78844
0.68978 | -33.9417 | -0.80555
0.69589 | -33.9409 | -0.82445
0.70199 | -33.9401 | -0.84677
0.70809 | -33.9393 | -0.86576
0.7142 | -33.9387 | -0.88598
0.7203 | -33.9378 | -0.90497
0.72641 | -33.9373 | -0.9271
0.73251 | -33.9367 | -0.94691
0.73862 | -33.937 | -0.96838
0.74472 | -33.9368 | -0.98821
0.75082 | -33.9367 | -1.01
0.75693 | -33.937 | -1.0315
0.76303 | -33.9376 | -1.0532
0.76914 | -33.9394 | -1.0735
0.77524 | -33.9404 | -1.093
0.78135 | -33.9417 | -1.1129
0.78745 | -33.943 | -1.1326
0.79355 | -33.9442 | -1.1522
0.79966 | -33.9459 | -1.1715
0.80576 | -33.9481 | -1.1909
0.81187 | -33.9497 | -1.2122
0.81797 | -33.9516 | -1.2331
0.82408 | -33.9532 | -1.2534
0.83018 | -33.9549 | -1.2737
0.83628 | -33.9568 | -1.292
0.84239 | -33.9579 | -1.3121
0.84849 | -33.9593 | -1.3334
0.8546 | -33.9607 | -1.3529
0.8607 | -33.962 | -1.3732
0.86681 | -33.963 | -1.3919
0.87291 | -33.964 | -1.4104
0.87901 | -33.9645 | -1.4282
0.88512 | -33.9652 | -1.4461
0.89122 | -33.966 | -1.4641
0.89733 | -33.9664 | -1.4797
0.90343 | -33.9668 | -1.4971
0.90953 | -33.967 | -1.5114
0.91564 | -33.9671 | -1.5272
0.92174 | -33.9673 | -1.5422
0.92785 | -33.9675 | -1.5568
0.93395 | -33.9676 | -1.5707
0.94006 | -33.9678 | -1.5836
0.94616 | -33.9676 | -1.5957
0.95226 | -33.9677 | -1.6085
0.95837 | -33.9678 | -1.6224
0.96447 | -33.9681 | -1.6353
0.97058 | -33.9679 | -1.6488
0.97668 | -33.9678 | -1.6622
0.98279 | -33.9676 | -1.675
0.98889 | -33.9674 | -1.6859
0.99499 | -33.9674 | -1.6994
1.0011 | -33.9674 | -1.7103
1.0072 | -33.9671 | -1.7221
1.0133 | -33.9668 | -1.7331
1.0194 | -33.9665 | -1.7462
1.0255 | -33.9662 | -1.7599
1.0316 | -33.9657 | -1.772
1.0377 | -33.9654 | -1.7857
1.0438 | -33.9649 | -1.7985
1.0499 | -33.9643 | -1.8137
1.056 | -33.9637 | -1.828
1.0621 | -33.9631 | -1.8423
1.0682 | -33.9628 | -1.8548
1.0743 | -33.9622 | -1.8702
1.0805 | -33.9618 | -1.8846
1.0866 | -33.961 | -1.8996
1.0927 | -33.9603 | -1.9141
1.0988 | -33.9597 | -1.929
1.1049 | -33.959 | -1.9432
1.111 | -33.9584 | -1.9585
1.1171 | -33.9578 | -1.973
1.1232 | -33.9574 | -1.9881
1.1293 | -33.957 | -2.0022
1.1354 | -33.9568 | -2.0171
1.1415 | -33.9565 | -2.0326
1.1476 | -33.956 | -2.0479
1.1537 | -33.9556 | -2.0626
1.1598 | -33.9557 | -2.0774
1.1659 | -33.9555 | -2.0907
1.172 | -33.9553 | -2.1047
1.1781 | -33.9551 | -2.1186
1.1842 | -33.955 | -2.1327
1.1903 | -33.9554 | -2.1438
1.1964 | -33.9553 | -2.1572
1.2025 | -33.9554 | -2.1711
1.2086 | -33.9557 | -2.1836
1.2147 | -33.9558 | -2.1962
1.2209 | -33.9561 | -2.2075
1.227 | -33.9568 | -2.2184
1.2331 | -33.9572 | -2.2305
1.2392 | -33.9576 | -2.2423
1.2453 | -33.958 | -2.2535
1.2514 | -33.9584 | -2.2636
1.2575 | -33.9589 | -2.2745
1.2636 | -33.9597 | -2.2835
1.2697 | -33.9603 | -2.293
1.2758 | -33.9609 | -2.3041
1.2819 | -33.9616 | -2.3141
1.288 | -33.9622 | -2.3249
1.2941 | -33.963 | -2.3336
1.3002 | -33.9636 | -2.3418
1.3063 | -33.9642 | -2.351
1.3124 | -33.9648 | -2.3601
1.3185 | -33.9655 | -2.3692
1.3246 | -33.9661 | -2.3784
1.3307 | -33.967 | -2.3871
1.3368 | -33.9677 | -2.3959
1.3429 | -33.9683 | -2.4049
1.349 | -33.9688 | -2.4143
1.3551 | -33.9695 | -2.4237
1.3613 | -33.97 | -2.4321
1.3674 | -33.9706 | -2.4421
1.3735 | -33.9711 | -2.4519
1.3796 | -33.9715 | -2.4611
1.3857 | -33.9719 | -2.4704
1.3918 | -33.9724 | -2.4807
1.3979 | -33.973 | -2.49
1.404 | -33.9735 | -2.5013
1.4101 | -33.9739 | -2.5117
1.4162 | -33.9743 | -2.5223
1.4223 | -33.9748 | -2.5337
1.4284 | -33.9753 | -2.5438
1.4345 | -33.9755 | -2.5553
1.4406 | -33.9759 | -2.5688
1.4467 | -33.9763 | -2.5804
1.4528 | -33.9768 | -2.5917
1.4589 | -33.9773 | -2.6036
1.465 | -33.9775 | -2.6145
1.4711 | -33.978 | -2.6259
1.4772 | -33.9786 | -2.6392
1.4833 | -33.979 | -2.6521
1.4894 | -33.9796 | -2.6643
1.4955 | -33.9801 | -2.6762
1.5016 | -33.9807 | -2.6887
1.5078 | -33.9811 | -2.7025
1.5139 | -33.9815 | -2.7155
1.52 | -33.9819 | -2.7281
1.5261 | -33.9823 | -2.7407
1.5322 | -33.9826 | -2.7546
1.5383 | -33.983 | -2.7672
1.5444 | -33.9836 | -2.7798
1.5505 | -33.984 | -2.7932
1.5566 | -33.9846 | -2.8062
1.5627 | -33.985 | -2.8199
1.5688 | -33.9856 | -2.8334
1.5749 | -33.986 | -2.846
1.581 | -33.9863 | -2.8594
1.5871 | -33.9867 | -2.8731
1.5932 | -33.9871 | -2.8859
1.5993 | -33.9875 | -2.9001
1.6054 | -33.9879 | -2.9133
1.6115 | -33.9883 | -2.9273
1.6176 | -33.9886 | -2.9399
1.6237 | -33.9889 | -2.9536
1.6298 | -33.9892 | -2.9662
1.6359 | -33.9895 | -2.979
1.642 | -33.99 | -2.9911
1.6482 | -33.9902 | -3.0042
1.6543 | -33.9905 | -3.0178
1.6604 | -33.9909 | -3.0299
1.6665 | -33.991 | -3.0424
1.6726 | -33.9912 | -3.0551
1.6787 | -33.9916 | -3.0677
1.6848 | -33.9915 | -3.0804
1.6909 | -33.9918 | -3.0933
1.697 | -33.992 | -3.1056
1.7031 | -33.9921 | -3.1183
1.7092 | -33.9921 | -3.1306
1.7153 | -33.9923 | -3.1427
1.7214 | -33.9921 | -3.155
1.7275 | -33.9921 | -3.1677
1.7336 | -33.9921 | -3.1807
1.7397 | -33.9922 | -3.1931
1.7458 | -33.9922 | -3.2055
1.7519 | -33.9923 | -3.2178
1.758 | -33.9924 | -3.2298
1.7641 | -33.9925 | -3.2415
1.7702 | -33.9925 | -3.2546
1.7763 | -33.9925 | -3.2662
1.7824 | -33.9925 | -3.2788
1.7885 | -33.9925 | -3.2913
1.7947 | -33.9926 | -3.3034
1.8008 | -33.9927 | -3.3148
1.8069 | -33.9927 | -3.3266
1.813 | -33.9928 | -3.3388
1.8191 | -33.9928 | -3.3516
1.8252 | -33.9929 | -3.3643
1.8313 | -33.9931 | -3.3766
1.8374 | -33.9932 | -3.3885
1.8435 | -33.9933 | -3.4005
1.8496 | -33.9934 | -3.4135
1.8557 | -33.9936 | -3.4262
1.8618 | -33.9939 | -3.4383
1.8679 | -33.994 | -3.4504
1.874 | -33.9941 | -3.4624
1.8801 | -33.9943 | -3.4747
1.8862 | -33.9946 | -3.4867
1.8923 | -33.9947 | -3.4988
1.8984 | -33.9949 | -3.5108
1.9045 | -33.995 | -3.5224
1.9106 | -33.995 | -3.5345
1.9167 | -33.9951 | -3.5467
1.9228 | -33.9953 | -3.5585
1.9289 | -33.9954 | -3.5703
1.9351 | -33.9956 | -3.5811
1.9412 | -33.9959 | -3.5924
1.9473 | -33.9962 | -3.6044
1.9534 | -33.9964 | -3.6158
1.9595 | -33.9966 | -3.6276
1.9656 | -33.9968 | -3.6397
1.9717 | -33.9971 | -3.6513
1.9778 | -33.9974 | -3.6627
1.9839 | -33.9977 | -3.6741
1.99 | -33.9979 | -3.6851
1.9961 | -33.998 | -3.6972
2.0022 | -33.9983 | -3.7085
2.0083 | -33.9984 | -3.7197
2.0144 | -33.9986 | -3.7311
2.0205 | -33.9989 | -3.7428
2.0266 | -33.9993 | -3.7542
2.0327 | -33.9995 | -3.7656
2.0388 | -33.9999 | -3.7772
2.0449 | -34.0001 | -3.7885
2.051 | -34.0004 | -3.7998
2.0571 | -34.0008 | -3.8111
2.0632 | -34.001 | -3.8226
2.0693 | -34.0012 | -3.8347
2.0754 | -34.0014 | -3.846
2.0816 | -34.0017 | -3.8581
2.0877 | -34.0019 | -3.8696
2.0938 | -34.0023 | -3.8817
2.0999 | -34.0025 | -3.8934
2.106 | -34.0027 | -3.905
2.1121 | -34.0029 | -3.917
2.1182 | -34.0031 | -3.9278
2.1243 | -34.0034 | -3.9386
2.1304 | -34.0036 | -3.9497
2.1365 | -34.0039 | -3.9614
2.1426 | -34.0041 | -3.9734
2.1487 | -34.0044 | -3.9847
2.1548 | -34.0048 | -3.9965
2.1609 | -34.0052 | -4.0079
2.167 | -34.0056 | -4.0192
2.1731 | -34.0059 | -4.0311
2.1792 | -34.0061 | -4.0423
2.1853 | -34.0065 | -4.0542
2.1914 | -34.0068 | -4.0655
2.1975 | -34.0072 | -4.0771
2.2036 | -34.0074 | -4.0891
2.2097 | -34.0076 | -4.1006
2.2158 | -34.008 | -4.1113
2.222 | -34.0082 | -4.1219
2.2281 | -34.0085 | -4.1333
2.2342 | -34.0088 | -4.1446
2.2403 | -34.0089 | -4.1554
2.2464 | -34.0092 | -4.1662
2.2525 | -34.0095 | -4.1777
2.2586 | -34.0098 | -4.1883
2.2647 | -34.0101 | -4.1989
2.2708 | -34.0104 | -4.2102
2.2769 | -34.0107 | -4.2209
2.283 | -34.0109 | -4.2321
2.2891 | -34.0111 | -4.2431
2.2952 | -34.0112 | -4.2537
2.3013 | -34.0115 | -4.2647
2.3074 | -34.0117 | -4.2752
2.3135 | -34.0119 | -4.2866
2.3196 | -34.0121 | -4.2969
2.3257 | -34.0123 | -4.3076
2.3318 | -34.0125 | -4.3187
2.3379 | -34.0127 | -4.3296
2.344 | -34.0128 | -4.341
2.3501 | -34.013 | -4.3517
2.3562 | -34.0132 | -4.3622
2.3623 | -34.0133 | -4.3727
2.3685 | -34.0134 | -4.3833
2.3746 | -34.0135 | -4.3946
2.3807 | -34.0137 | -4.4052
2.3868 | -34.0139 | -4.416
2.3929 | -34.0141 | -4.4267
2.399 | -34.0142 | -4.4372
2.4051 | -34.0145 | -4.4478
2.4112 | -34.0146 | -4.4585
2.4173 | -34.0148 | -4.4692
2.4234 | -34.0149 | -4.4801
2.4295 | -34.0151 | -4.491
2.4356 | -34.0152 | -4.5019
2.4417 | -34.0154 | -4.513
2.4478 | -34.0155 | -4.524
2.4539 | -34.0156 | -4.5343
2.46 | -34.0157 | -4.5457
2.4661 | -34.0159 | -4.5569
2.4722 | -34.0161 | -4.5672
2.4783 | -34.0162 | -4.5777
2.4844 | -34.0163 | -4.5888
2.4905 | -34.0164 | -4.5998
2.4966 | -34.0166 | -4.6105
2.5027 | -34.0167 | -4.6211
2.5089 | -34.0168 | -4.632
2.515 | -34.0168 | -4.6429
2.5211 | -34.017 | -4.6537
2.5272 | -34.0171 | -4.6647
2.5333 | -34.0172 | -4.6757
2.5394 | -34.0174 | -4.6865
2.5455 | -34.0175 | -4.6978
2.5516 | -34.0177 | -4.7083
2.5577 | -34.0179 | -4.7197
2.5638 | -34.018 | -4.7315
2.5699 | -34.0181 | -4.7429
2.576 | -34.0182 | -4.7545
2.5821 | -34.0183 | -4.7658
2.5882 | -34.0184 | -4.7769
2.5943 | -34.0186 | -4.7878
2.6004 | -34.0188 | -4.7989
2.6065 | -34.0189 | -4.8093
2.6126 | -34.0191 | -4.8199
2.6187 | -34.0192 | -4.8305
2.6248 | -34.0194 | -4.8414
2.6309 | -34.0195 | -4.8528
2.637 | -34.0198 | -4.8645
2.6431 | -34.02 | -4.8751
2.6492 | -34.0201 | -4.8865
2.6554 | -34.0203 | -4.8974
2.6615 | -34.0204 | -4.9085
2.6676 | -34.0207 | -4.9199
2.6737 | -34.021 | -4.9311
2.6798 | -34.0213 | -4.9423
2.6859 | -34.0215 | -4.9537
2.692 | -34.0217 | -4.9646
2.6981 | -34.0219 | -4.9758
2.7042 | -34.0221 | -4.9869
2.7103 | -34.0223 | -4.9976
2.7164 | -34.0226 | -5.0087
2.7225 | -34.0229 | -5.0201
2.7286 | -34.0231 | -5.0311
2.7347 | -34.0233 | -5.0418
2.7408 | -34.0236 | -5.052
2.7469 | -34.0239 | -5.0633
2.753 | -34.0242 | -5.074
2.7591 | -34.0245 | -5.0853
2.7652 | -34.0247 | -5.0969
2.7713 | -34.0249 | -5.1077
2.7774 | -34.0252 | -5.1187
2.7835 | -34.0254 | -5.1298
2.7896 | -34.0256 | -5.1404
2.7958 | -34.0259 | -5.1512
2.8019 | -34.0262 | -5.1617
2.808 | -34.0264 | -5.1726
2.8141 | -34.0268 | -5.1839
2.8202 | -34.0271 | -5.1948
2.8263 | -34.0274 | -5.2058
2.8324 | -34.0277 | -5.2172
2.8385 | -34.0281 | -5.2281
2.8446 | -34.0285 | -5.2389
2.8507 | -34.0289 | -5.2498
2.8568 | -34.0292 | -5.2604
2.8629 | -34.0295 | -5.271
2.869 | -34.0298 | -5.2822
2.8751 | -34.0301 | -5.293
2.8812 | -34.0305 | -5.3043
2.8873 | -34.0308 | -5.3152
2.8934 | -34.0312 | -5.3263
2.8995 | -34.0315 | -5.3374
2.9056 | -34.0318 | -5.3486
2.9117 | -34.0321 | -5.3597
2.9178 | -34.0325 | -5.3705
2.9239 | -34.0329 | -5.382
2.93 | -34.0333 | -5.3935
2.9361 | -34.0336 | -5.4043
2.9423 | -34.0339 | -5.415
2.9484 | -34.0343 | -5.4262
2.9545 | -34.0347 | -5.4372
2.9606 | -34.0351 | -5.4486
2.9667 | -34.0354 | -5.4597
2.9728 | -34.0358 | -5.471
2.9789 | -34.0362 | -5.4821
2.985 | -34.0366 | -5.4933
2.9911 | -34.037 | -5.5046
2.9972 | -34.0374 | -5.5157
3.0033 | -34.0378 | -5.5272
3.0094 | -34.0382 | -5.5383
3.0155 | -34.0386 | -5.5492
3.0216 | -34.0389 | -5.5605
3.0277 | -34.0393 | -5.5722
3.0338 | -34.0397 | -5.5835
3.0399 | -34.0401 | -5.5948
3.046 | -34.0404 | -5.6063
3.0521 | -34.0407 | -5.6174
3.0582 | -34.041 | -5.6293
3.0643 | -34.0413 | -5.6407
3.0704 | -34.0416 | -5.6525
3.0765 | -34.0419 | -5.6636
3.0827 | -34.0422 | -5.6748
3.0888 | -34.0425 | -5.686
3.0949 | -34.0429 | -5.6973
3.101 | -34.0432 | -5.7087
3.1071 | -34.0436 | -5.7204
3.1132 | -34.0439 | -5.732
3.1193 | -34.0443 | -5.7434
3.1254 | -34.0447 | -5.7547
3.1315 | -34.045 | -5.7665
3.1376 | -34.0453 | -5.7779
3.1437 | -34.0456 | -5.7894
3.1498 | -34.046 | -5.8014
3.1559 | -34.0463 | -5.8128
3.162 | -34.0465 | -5.8247
3.1681 | -34.0468 | -5.8366
3.1742 | -34.0471 | -5.8481
3.1803 | -34.0474 | -5.8597
3.1864 | -34.0478 | -5.8709
3.1925 | -34.0481 | -5.8824
3.1986 | -34.0483 | -5.8943
3.2047 | -34.0486 | -5.9062
3.2108 | -34.0488 | -5.9173
3.2169 | -34.0492 | -5.9289
3.223 | -34.0494 | -5.9405
3.2292 | -34.0496 | -5.9523
3.2353 | -34.0498 | -5.9631
3.2414 | -34.0501 | -5.9745
3.2475 | -34.0503 | -5.9859
3.2536 | -34.0506 | -5.9972
3.2597 | -34.0508 | -6.0082
3.2658 | -34.051 | -6.0191
3.2719 | -34.0513 | -6.0302
3.278 | -34.0515 | -6.0417
3.2841 | -34.0518 | -6.0528
3.2902 | -34.052 | -6.0642
3.2963 | -34.0523 | -6.0756
3.3024 | -34.0524 | -6.0867
3.3085 | -34.0527 | -6.0982
3.3146 | -34.0528 | -6.1094
3.3207 | -34.0531 | -6.1209
3.3268 | -34.0533 | -6.1323
3.3329 | -34.0535 | -6.143
3.339 | -34.0537 | -6.1539
3.3451 | -34.0539 | -6.1651
3.3512 | -34.0541 | -6.1758
3.3573 | -34.0543 | -6.187
3.3634 | -34.0545 | -6.1982
3.3696 | -34.0547 | -6.2096
3.3757 | -34.0549 | -6.2204
3.3818 | -34.0551 | -6.2313
3.3879 | -34.0553 | -6.2424
3.394 | -34.0554 | -6.2538
3.4001 | -34.0556 | -6.2647
3.4062 | -34.0558 | -6.2757
3.4123 | -34.0559 | -6.2863
3.4184 | -34.0561 | -6.2976
3.4245 | -34.0564 | -6.3085
3.4306 | -34.0565 | -6.3192
3.4367 | -34.0567 | -6.3298
3.4428 | -34.0569 | -6.3404
3.4489 | -34.057 | -6.3513
3.455 | -34.0572 | -6.3623
3.4611 | -34.0574 | -6.3735
3.4672 | -34.0575 | -6.3847
3.4733 | -34.0577 | -6.3958
3.4794 | -34.0579 | -6.4067
3.4855 | -34.0581 | -6.4172
3.4916 | -34.0583 | -6.4277
3.4977 | -34.0584 | -6.4391
3.5038 | -34.0586 | -6.45
3.5099 | -34.0587 | -6.4604
3.5161 | -34.0589 | -6.4709
3.5222 | -34.059 | -6.4814
3.5283 | -34.0592 | -6.4923
3.5344 | -34.0593 | -6.5033
3.5405 | -34.0595 | -6.5139
3.5466 | -34.0597 | -6.5249
3.5527 | -34.0598 | -6.5359
3.5588 | -34.0599 | -6.5466
3.5649 | -34.06 | -6.5575
3.571 | -34.0602 | -6.5682
3.5771 | -34.0603 | -6.5793
3.5832 | -34.0605 | -6.5894
3.5893 | -34.0607 | -6.6001
3.5954 | -34.0609 | -6.6111
3.6015 | -34.0611 | -6.6221
3.6076 | -34.0613 | -6.6327
3.6137 | -34.0614 | -6.643
3.6198 | -34.0617 | -6.6534
3.6259 | -34.0619 | -6.6644
3.632 | -34.0621 | -6.6752
3.6381 | -34.0622 | -6.6861
3.6442 | -34.0624 | -6.6968
3.6503 | -34.0626 | -6.7071
3.6565 | -34.0629 | -6.718
3.6626 | -34.063 | -6.7287
3.6687 | -34.0632 | -6.7395
3.6748 | -34.0635 | -6.75
3.6809 | -34.0638 | -6.7608
3.687 | -34.0639 | -6.7719
3.6931 | -34.0642 | -6.7822
3.6992 | -34.0644 | -6.793
3.7053 | -34.0646 | -6.8037
3.7114 | -34.0648 | -6.8145
3.7175 | -34.065 | -6.8253
3.7236 | -34.0653 | -6.8355
3.7297 | -34.0655 | -6.8464
3.7358 | -34.0657 | -6.8574
3.7419 | -34.0659 | -6.8683
3.748 | -34.0661 | -6.879
3.7541 | -34.0664 | -6.8897
3.7602 | -34.0666 | -6.9004
3.7663 | -34.0668 | -6.911
3.7724 | -34.0671 | -6.9218
3.7785 | -34.0674 | -6.9323
3.7846 | -34.0676 | -6.9429
3.7907 | -34.0679 | -6.9539
3.7969 | -34.0681 | -6.9648
3.803 | -34.0684 | -6.9751
3.8091 | -34.0686 | -6.9858
3.8152 | -34.0689 | -6.9964
3.8213 | -34.0691 | -7.0074
3.8274 | -34.0694 | -7.0178
3.8335 | -34.0697 | -7.0287
3.8396 | -34.07 | -7.0395
3.8457 | -34.0703 | -7.0503
3.8518 | -34.0706 | -7.0609
3.8579 | -34.071 | -7.0716
3.864 | -34.0712 | -7.0825
3.8701 | -34.0715 | -7.0936
3.8762 | -34.0718 | -7.1041
3.8823 | -34.072 | -7.115
3.8884 | -34.0724 | -7.1258
3.8945 | -34.0727 | -7.1368
3.9006 | -34.073 | -7.1478
3.9067 | -34.0732 | -7.159
3.9128 | -34.0735 | -7.1698
3.9189 | -34.0738 | -7.1804
3.925 | -34.074 | -7.1913
3.9311 | -34.0743 | -7.2017
3.9372 | -34.0746 | -7.2126
3.9434 | -34.0749 | -7.2234
3.9495 | -34.0753 | -7.2342
3.9556 | -34.0756 | -7.2448
3.9617 | -34.0759 | -7.2552
3.9678 | -34.0762 | -7.2657
3.9739 | -34.0764 | -7.2764
3.98 | -34.0767 | -7.2873
3.9861 | -34.077 | -7.298
3.9922 | -34.0772 | -7.3088
3.9983 | -34.0775 | -7.3192
4.0044 | -34.0778 | -7.3296
4.0105 | -34.0781 | -7.3403
4.0166 | -34.0784 | -7.3511
4.0227 | -34.0787 | -7.3618
4.0288 | -34.0789 | -7.3729
4.0349 | -34.0792 | -7.3835
4.041 | -34.0794 | -7.3942
4.0471 | -34.0796 | -7.405
4.0532 | -34.0799 | -7.4154
4.0593 | -34.0802 | -7.4263
4.0654 | -34.0805 | -7.4368
4.0715 | -34.0808 | -7.448
4.0776 | -34.0811 | -7.4586
4.0838 | -34.0814 | -7.4697
4.0899 | -34.0817 | -7.4802
4.096 | -34.0819 | -7.4909
4.1021 | -34.0822 | -7.5014
4.1082 | -34.0824 | -7.5119
4.1143 | -34.0827 | -7.5223
4.1204 | -34.083 | -7.5327
4.1265 | -34.0832 | -7.5431
4.1326 | -34.0835 | -7.5535
4.1387 | -34.0838 | -7.5641
4.1448 | -34.0841 | -7.5749
4.1509 | -34.0843 | -7.5855
4.157 | -34.0845 | -7.5966
4.1631 | -34.0848 | -7.6074
4.1692 | -34.0851 | -7.618
4.1753 | -34.0853 | -7.6286
4.1814 | -34.0855 | -7.6389
4.1875 | -34.0858 | -7.6498
4.1936 | -34.0861 | -7.6605
4.1997 | -34.0863 | -7.6708
4.2058 | -34.0866 | -7.6816
4.2119 | -34.0868 | -7.6921
4.218 | -34.087 | -7.7024
4.2241 | -34.0872 | -7.7128
4.2303 | -34.0875 | -7.7233
4.2364 | -34.0878 | -7.7339
4.2425 | -34.0879 | -7.7445
4.2486 | -34.0882 | -7.7552
4.2547 | -34.0883 | -7.7656
4.2608 | -34.0885 | -7.7763
4.2669 | -34.0887 | -7.7865
4.273 | -34.089 | -7.7972
4.2791 | -34.0892 | -7.8077
4.2852 | -34.0894 | -7.8185
4.2913 | -34.0896 | -7.8288
4.2974 | -34.0898 | -7.8396
4.3035 | -34.0901 | -7.8504
4.3096 | -34.0903 | -7.8611
4.3157 | -34.0905 | -7.8715
4.3218 | -34.0907 | -7.8819
4.3279 | -34.0909 | -7.8925
4.334 | -34.0911 | -7.9029
4.3401 | -34.0913 | -7.9131
4.3462 | -34.0916 | -7.9237
4.3523 | -34.0919 | -7.9338
4.3584 | -34.092 | -7.944
4.3645 | -34.0922 | -7.9544
4.3707 | -34.0925 | -7.9646
4.3768 | -34.0927 | -7.975
4.3829 | -34.0929 | -7.9858
4.389 | -34.0931 | -7.9963
4.3951 | -34.0933 | -8.0065
4.4012 | -34.0935 | -8.0173
4.4073 | -34.0937 | -8.0275
4.4134 | -34.0939 | -8.0379
4.4195 | -34.0942 | -8.0483
4.4256 | -34.0945 | -8.0586
4.4317 | -34.0947 | -8.0689
4.4378 | -34.0949 | -8.079
4.4439 | -34.0951 | -8.0891
4.45 | -34.0952 | -8.0995
4.4561 | -34.0954 | -8.1098
4.4622 | -34.0956 | -8.1202
4.4683 | -34.0958 | -8.1306
4.4744 | -34.096 | -8.1413
4.4805 | -34.0962 | -8.1512
4.4866 | -34.0964 | -8.1616
4.4927 | -34.0967 | -8.1722
4.4988 | -34.0969 | -8.1825
4.5049 | -34.0971 | -8.1925
4.511 | -34.0974 | -8.2025
4.5172 | -34.0976 | -8.2129
4.5233 | -34.0978 | -8.2236
4.5294 | -34.098 | -8.2338
4.5355 | -34.0982 | -8.2441
4.5416 | -34.0985 | -8.2544
4.5477 | -34.0987 | -8.2647
4.5538 | -34.0989 | -8.2747
4.5599 | -34.0992 | -8.2849
4.566 | -34.0994 | -8.2953
4.5721 | -34.0996 | -8.3056
4.5782 | -34.0999 | -8.3161
4.5843 | -34.1001 | -8.3265
4.5904 | -34.1004 | -8.3369
4.5965 | -34.1006 | -8.3471
4.6026 | -34.1008 | -8.3578
4.6087 | -34.101 | -8.368
4.6148 | -34.1012 | -8.3786
4.6209 | -34.1015 | -8.3888
4.627 | -34.1018 | -8.3993
4.6331 | -34.1019 | -8.4096
4.6392 | -34.1022 | -8.4199
4.6453 | -34.1024 | -8.4303
4.6514 | -34.1027 | -8.4406
4.6576 | -34.1029 | -8.4513
4.6637 | -34.1031 | -8.4613
4.6698 | -34.1033 | -8.4715
4.6759 | -34.1035 | -8.4822
4.682 | -34.1037 | -8.4925
4.6881 | -34.1039 | -8.5028
4.6942 | -34.1041 | -8.513
4.7003 | -34.1044 | -8.5234
4.7064 | -34.1046 | -8.5336
4.7125 | -34.1049 | -8.5442
4.7186 | -34.1051 | -8.5544
4.7247 | -34.1053 | -8.5648
4.7308 | -34.1056 | -8.5751
4.7369 | -34.1058 | -8.5855
4.743 | -34.1061 | -8.5962
4.7491 | -34.1063 | -8.6068
4.7552 | -34.1065 | -8.6173
4.7613 | -34.1067 | -8.6278
4.7674 | -34.1069 | -8.638
4.7735 | -34.1072 | -8.6487
4.7796 | -34.1075 | -8.6593
4.7857 | -34.1077 | -8.6698
4.7918 | -34.108 | -8.68
4.7979 | -34.1082 | -8.6903
4.8041 | -34.1085 | -8.7007
4.8102 | -34.1087 | -8.7108
4.8163 | -34.109 | -8.7211
4.8224 | -34.1092 | -8.7315
4.8285 | -34.1095 | -8.7416
4.8346 | -34.1097 | -8.7518
4.8407 | -34.1099 | -8.7623
4.8468 | -34.1102 | -8.773
4.8529 | -34.1105 | -8.7836
4.859 | -34.1107 | -8.7943
4.8651 | -34.111 | -8.8045
4.8712 | -34.1112 | -8.8148
4.8773 | -34.1114 | -8.8253
4.8834 | -34.1117 | -8.8357
4.8895 | -34.1119 | -8.8459
4.8956 | -34.1121 | -8.8563
4.9017 | -34.1124 | -8.8666
4.9078 | -34.1126 | -8.8773
4.9139 | -34.1129 | -8.8873
4.92 | -34.1132 | -8.8974
4.9261 | -34.1135 | -8.9075
4.9322 | -34.1137 | -8.9179
4.9383 | -34.114 | -8.9283
4.9445 | -34.1143 | -8.9387
4.9506 | -34.1145 | -8.9489
4.9567 | -34.1148 | -8.9591
4.9628 | -34.115 | -8.9695
4.9689 | -34.1153 | -8.9803
4.975 | -34.1155 | -8.991
4.9811 | -34.1158 | -9.0011
4.9872 | -34.1161 | -9.0113
4.9933 | -34.1164 | -9.022
4.9994 | -34.1167 | -9.0328
5.0055 | -34.117 | -9.0432
5.0116 | -34.1172 | -9.054
5.0177 | -34.1176 | -9.0647
5.0238 | -34.1179 | -9.0754
5.0299 | -34.1182 | -9.0861
5.036 | -34.1185 | -9.0965
5.0421 | -34.1187 | -9.1069
5.0482 | -34.119 | -9.1173
5.0543 | -34.1193 | -9.1275
5.0604 | -34.1196 | -9.1379
5.0665 | -34.1199 | -9.1481
5.0726 | -34.1202 | -9.1589
5.0787 | -34.1205 | -9.1695
5.0848 | -34.1208 | -9.1801
5.091 | -34.1211 | -9.1903
5.0971 | -34.1214 | -9.2005
5.1032 | -34.1217 | -9.2113
5.1093 | -34.122 | -9.2217
5.1154 | -34.1222 | -9.232
5.1215 | -34.1225 | -9.2422
5.1276 | -34.1228 | -9.2525
5.1337 | -34.1231 | -9.2628
5.1398 | -34.1234 | -9.273
5.1459 | -34.1237 | -9.2837
5.152 | -34.124 | -9.2943
5.1581 | -34.1243 | -9.3049
5.1642 | -34.1246 | -9.3154
5.1703 | -34.1249 | -9.3258
5.1764 | -34.1252 | -9.3361
5.1825 | -34.1255 | -9.3464
5.1886 | -34.1258 | -9.3565
5.1947 | -34.126 | -9.367
5.2008 | -34.1264 | -9.3775
5.2069 | -34.1267 | -9.3882
5.213 | -34.127 | -9.3984
5.2191 | -34.1273 | -9.4089
5.2252 | -34.1276 | -9.4195
5.2314 | -34.1279 | -9.4301
5.2375 | -34.1282 | -9.4403
5.2436 | -34.1284 | -9.4508
5.2497 | -34.1287 | -9.461
5.2558 | -34.1291 | -9.4718
5.2619 | -34.1294 | -9.4825
5.268 | -34.1297 | -9.4931
5.2741 | -34.13 | -9.5037
5.2802 | -34.1303 | -9.5145
5.2863 | -34.1307 | -9.5248
5.2924 | -34.131 | -9.5355
5.2985 | -34.1313 | -9.546
5.3046 | -34.1316 | -9.5566
5.3107 | -34.1319 | -9.5672
5.3168 | -34.1322 | -9.5778
5.3229 | -34.1325 | -9.5883
5.329 | -34.1328 | -9.5988
5.3351 | -34.1332 | -9.609
5.3412 | -34.1335 | -9.6195
5.3473 | -34.1338 | -9.63
5.3534 | -34.1341 | -9.6406
5.3595 | -34.1344 | -9.6511
5.3656 | -34.1347 | -9.6616
5.3717 | -34.135 | -9.6723
5.3779 | -34.1353 | -9.6832
5.384 | -34.1356 | -9.6937
5.3901 | -34.136 | -9.7044
5.3962 | -34.1363 | -9.7147
5.4023 | -34.1366 | -9.7255
5.4084 | -34.137 | -9.7359
5.4145 | -34.1373 | -9.7461
5.4206 | -34.1376 | -9.7567
5.4267 | -34.1379 | -9.7675
5.4328 | -34.1382 | -9.778
5.4389 | -34.1385 | -9.7883
5.445 | -34.1388 | -9.7991
5.4511 | -34.1392 | -9.8095
5.4572 | -34.1395 | -9.8201
5.4633 | -34.1398 | -9.8306
5.4694 | -34.1401 | -9.8411
5.4755 | -34.1404 | -9.8514
5.4816 | -34.1407 | -9.862
5.4877 | -34.141 | -9.8724
5.4938 | -34.1413 | -9.8826
5.4999 | -34.1417 | -9.893
5.506 | -34.142 | -9.9035
5.5121 | -34.1423 | -9.9139
5.5183 | -34.1426 | -9.9242
5.5244 | -34.1429 | -9.9349
5.5305 | -34.1433 | -9.9453
5.5366 | -34.1436 | -9.9555
5.5427 | -34.1439 | -9.9662
5.5488 | -34.1442 | -9.9767
5.5549 | -34.1445 | -9.9872
5.561 | -34.1448 | -9.9979
5.5671 | -34.1451 | -10.0083
5.5732 | -34.1454 | -10.019
5.5793 | -34.1458 | -10.0298
5.5854 | -34.1461 | -10.0402
5.5915 | -34.1464 | -10.051
5.5976 | -34.1467 | -10.0615
5.6037 | -34.147 | -10.0721
5.6098 | -34.1473 | -10.0826
5.6159 | -34.1476 | -10.0929
5.622 | -34.1479 | -10.1034
5.6281 | -34.1483 | -10.114
5.6342 | -34.1486 | -10.1247
5.6403 | -34.1489 | -10.1353
5.6464 | -34.1492 | -10.1459
5.6525 | -34.1495 | -10.1566
5.6586 | -34.1498 | -10.1668
5.6648 | -34.1502 | -10.1774
5.6709 | -34.1505 | -10.188
5.677 | -34.1508 | -10.1986
5.6831 | -34.1511 | -10.2092
5.6892 | -34.1514 | -10.2198
5.6953 | -34.1518 | -10.2303
5.7014 | -34.1521 | -10.2408
5.7075 | -34.1523 | -10.2514
5.7136 | -34.1526 | -10.2618
5.7197 | -34.1529 | -10.2723
5.7258 | -34.1532 | -10.2826
5.7319 | -34.1536 | -10.2932
5.738 | -34.1539 | -10.3037
5.7441 | -34.1542 | -10.3141
5.7502 | -34.1546 | -10.3244
5.7563 | -34.1549 | -10.3349
5.7624 | -34.1552 | -10.3456
5.7685 | -34.1555 | -10.3561
5.7746 | -34.1558 | -10.3668
5.7807 | -34.1561 | -10.3774
5.7868 | -34.1564 | -10.3881
5.7929 | -34.1567 | -10.3985
5.799 | -34.157 | -10.4088
5.8052 | -34.1573 | -10.4192
5.8113 | -34.1576 | -10.4295
5.8174 | -34.1579 | -10.4399
5.8235 | -34.1582 | -10.4503
5.8296 | -34.1585 | -10.461
5.8357 | -34.1588 | -10.4714
5.8418 | -34.1591 | -10.482
5.8479 | -34.1594 | -10.4926
5.854 | -34.1597 | -10.5032
5.8601 | -34.16 | -10.5136
5.8662 | -34.1603 | -10.5242
5.8723 | -34.1606 | -10.5345
5.8784 | -34.1609 | -10.5449
5.8845 | -34.1612 | -10.5553
5.8906 | -34.1615 | -10.5658
5.8967 | -34.1618 | -10.5761
5.9028 | -34.1621 | -10.5864
5.9089 | -34.1625 | -10.5968
5.915 | -34.1628 | -10.6071
5.9211 | -34.1631 | -10.6174
5.9272 | -34.1634 | -10.6277
5.9333 | -34.1637 | -10.6379
5.9394 | -34.164 | -10.6484
5.9455 | -34.1643 | -10.6586
5.9517 | -34.1647 | -10.6689
5.9578 | -34.165 | -10.6794
5.9639 | -34.1654 | -10.6895
5.97 | -34.1657 | -10.7001
5.9761 | -34.166 | -10.7104
5.9822 | -34.1663 | -10.721
5.9883 | -34.1666 | -10.7314
5.9944 | -34.167 | -10.7419
6.0005 | -34.1673 | -10.7525
6.0066 | -34.1676 | -10.7628
6.0127 | -34.168 | -10.7729
6.0188 | -34.1683 | -10.7836
6.0249 | -34.1686 | -10.794
6.031 | -34.1689 | -10.8042
6.0371 | -34.1692 | -10.8147
6.0432 | -34.1695 | -10.825
6.0493 | -34.1698 | -10.8353
6.0554 | -34.1702 | -10.8455
6.0615 | -34.1705 | -10.8559
6.0676 | -34.1708 | -10.8665
6.0737 | -34.1711 | -10.877
6.0798 | -34.1714 | -10.8874
6.0859 | -34.1718 | -10.8977
6.0921 | -34.1721 | -10.9079
6.0982 | -34.1724 | -10.9182
6.1043 | -34.1728 | -10.9285
6.1104 | -34.1731 | -10.9391
6.1165 | -34.1734 | -10.9491
6.1226 | -34.1737 | -10.9594
6.1287 | -34.174 | -10.9697
6.1348 | -34.1744 | -10.98
6.1409 | -34.1747 | -10.9904
6.147 | -34.175 | -11.0007
6.1531 | -34.1753 | -11.011
6.1592 | -34.1757 | -11.0212
6.1653 | -34.176 | -11.0315
6.1714 | -34.1763 | -11.0419
6.1775 | -34.1766 | -11.0523
6.1836 | -34.1769 | -11.0627
6.1897 | -34.1773 | -11.0731
6.1958 | -34.1776 | -11.0836
6.2019 | -34.1779 | -11.094
6.208 | -34.1783 | -11.1044
6.2141 | -34.1786 | -11.1147
6.2202 | -34.1789 | -11.1251
6.2263 | -34.1793 | -11.1354
6.2325 | -34.1796 | -11.1455
6.2386 | -34.1799 | -11.1555
6.2447 | -34.1802 | -11.1661
6.2508 | -34.1806 | -11.1764
6.2569 | -34.1809 | -11.1867
6.263 | -34.1812 | -11.197
6.2691 | -34.1815 | -11.2072
6.2752 | -34.1818 | -11.2176
6.2813 | -34.1821 | -11.2279
6.2874 | -34.1825 | -11.2381
6.2935 | -34.1828 | -11.2484
6.2996 | -34.1831 | -11.2586
6.3057 | -34.1834 | -11.269
6.3118 | -34.1837 | -11.2794
6.3179 | -34.184 | -11.2897
6.324 | -34.1844 | -11.3
6.3301 | -34.1847 | -11.3104
6.3362 | -34.185 | -11.3206
6.3423 | -34.1854 | -11.331
6.3484 | -34.1857 | -11.3414
6.3545 | -34.186 | -11.3515
6.3606 | -34.1863 | -11.3615
6.3667 | -34.1866 | -11.3717
6.3728 | -34.1869 | -11.382
6.379 | -34.1873 | -11.3923
6.3851 | -34.1876 | -11.4024
6.3912 | -34.1879 | -11.4127
6.3973 | -34.1882 | -11.423
6.4034 | -34.1885 | -11.4331
6.4095 | -34.1888 | -11.4434
6.4156 | -34.1891 | -11.4537
6.4217 | -34.1894 | -11.4639
6.4278 | -34.1898 | -11.4742
6.4339 | -34.1901 | -11.4844
6.44 | -34.1905 | -11.4946
6.4461 | -34.1908 | -11.505
6.4522 | -34.1911 | -11.5151
6.4583 | -34.1914 | -11.5255
6.4644 | -34.1917 | -11.5358
6.4705 | -34.192 | -11.5458
6.4766 | -34.1923 | -11.5561
6.4827 | -34.1927 | -11.5664
6.4888 | -34.1929 | -11.5765
6.4949 | -34.1933 | -11.5865
6.501 | -34.1936 | -11.5969
6.5071 | -34.1939 | -11.607
6.5132 | -34.1943 | -11.6172
6.5194 | -34.1946 | -11.6275
6.5255 | -34.1949 | -11.6375
6.5316 | -34.1952 | -11.6479
6.5377 | -34.1955 | -11.6581
6.5438 | -34.1958 | -11.6682
6.5499 | -34.1961 | -11.6784
6.556 | -34.1965 | -11.6886
6.5621 | -34.1967 | -11.699
6.5682 | -34.1971 | -11.7093
6.5743 | -34.1974 | -11.7195
6.5804 | -34.1977 | -11.7298
6.5865 | -34.1981 | -11.7401
6.5926 | -34.1984 | -11.7502
6.5987 | -34.1987 | -11.7602
6.6048 | -34.199 | -11.7705
6.6109 | -34.1994 | -11.7807
6.617 | -34.1997 | -11.7911
6.6231 | -34.2 | -11.8013
6.6292 | -34.2003 | -11.8115
6.6353 | -34.2006 | -11.822
6.6414 | -34.201 | -11.8321
6.6475 | -34.2013 | -11.8424
6.6536 | -34.2016 | -11.8526
6.6597 | -34.2019 | -11.8626
6.6659 | -34.2022 | -11.8728
6.672 | -34.2026 | -11.8829
6.6781 | -34.2029 | -11.8932
6.6842 | -34.2032 | -11.9032
6.6903 | -34.2036 | -11.9134
6.6964 | -34.2039 | -11.9235
6.7025 | -34.2042 | -11.9336
6.7086 | -34.2045 | -11.9438
6.7147 | -34.2049 | -11.954
6.7208 | -34.2052 | -11.9643
6.7269 | -34.2055 | -11.9745
6.733 | -34.2058 | -11.9844
6.7391 | -34.2061 | -11.9946
6.7452 | -34.2065 | -12.0046
6.7513 | -34.2068 | -12.0146
6.7574 | -34.2071 | -12.0248
6.7635 | -34.2074 | -12.0348
6.7696 | -34.2078 | -12.0449
6.7757 | -34.2081 | -12.055
6.7818 | -34.2085 | -12.0652
6.7879 | -34.2088 | -12.0753
6.794 | -34.2091 | -12.0853
6.8001 | -34.2094 | -12.0954
6.8063 | -34.2098 | -12.1056
6.8124 | -34.2101 | -12.1159
6.8185 | -34.2104 | -12.1262
6.8246 | -34.2107 | -12.1363
6.8307 | -34.211 | -12.1463
6.8368 | -34.2113 | -12.1567
6.8429 | -34.2117 | -12.1669
6.849 | -34.212 | -12.1771
6.8551 | -34.2124 | -12.1871
6.8612 | -34.2127 | -12.1972
6.8673 | -34.213 | -12.2073
6.8734 | -34.2134 | -12.2176
6.8795 | -34.2137 | -12.2278
6.8856 | -34.214 | -12.2379
6.8917 | -34.2143 | -12.248
6.8978 | -34.2147 | -12.258
6.9039 | -34.215 | -12.2678
6.91 | -34.2153 | -12.278
6.9161 | -34.2157 | -12.2879
6.9222 | -34.216 | -12.298
6.9283 | -34.2164 | -12.3079
6.9344 | -34.2167 | -12.318
6.9405 | -34.2171 | -12.328
6.9466 | -34.2174 | -12.3383
6.9528 | -34.2178 | -12.3486
6.9589 | -34.2181 | -12.3588
6.965 | -34.2184 | -12.369
6.9711 | -34.2188 | -12.3792
6.9772 | -34.2192 | -12.3893
6.9833 | -34.2195 | -12.3994
6.9894 | -34.2198 | -12.4097
6.9955 | -34.2202 | -12.4196
7.0016 | -34.2205 | -12.4297
7.0077 | -34.2209 | -12.4397
7.0138 | -34.2212 | -12.4498
7.0199 | -34.2215 | -12.4598
7.026 | -34.2219 | -12.4699
7.0321 | -34.2222 | -12.4797
7.0382 | -34.2225 | -12.4896
7.0443 | -34.2229 | -12.4998
7.0504 | -34.2232 | -12.5099
7.0565 | -34.2236 | -12.52
7.0626 | -34.2239 | -12.5298
7.0687 | -34.2243 | -12.5395
7.0748 | -34.2246 | -12.5494
7.0809 | -34.225 | -12.5594
7.087 | -34.2253 | -12.5696
7.0932 | -34.2256 | -12.5797
7.0993 | -34.226 | -12.5897
7.1054 | -34.2263 | -12.5996
7.1115 | -34.2267 | -12.6095
7.1176 | -34.227 | -12.6194
7.1237 | -34.2273 | -12.6294
7.1298 | -34.2277 | -12.6394
7.1359 | -34.228 | -12.6494
7.142 | -34.2284 | -12.6592
7.1481 | -34.2287 | -12.6691
7.1542 | -34.229 | -12.6791
7.1603 | -34.2294 | -12.6891
7.1664 | -34.2297 | -12.6991
7.1725 | -34.2301 | -12.7091
7.1786 | -34.2305 | -12.719
7.1847 | -34.2308 | -12.7291
7.1908 | -34.2311 | -12.7393
7.1969 | -34.2315 | -12.7494
7.203 | -34.2318 | -12.7592
7.2091 | -34.2322 | -12.769
7.2152 | -34.2325 | -12.779
7.2213 | -34.2329 | -12.7888
7.2274 | -34.2332 | -12.7988
7.2335 | -34.2336 | -12.8087
7.2397 | -34.2339 | -12.8186
7.2458 | -34.2343 | -12.8283
7.2519 | -34.2347 | -12.8383
7.258 | -34.235 | -12.8484
7.2641 | -34.2354 | -12.8584
7.2702 | -34.2357 | -12.8685
7.2763 | -34.2361 | -12.8785
7.2824 | -34.2364 | -12.8885
7.2885 | -34.2368 | -12.8985
7.2946 | -34.2371 | -12.9084
7.3007 | -34.2374 | -12.9184
7.3068 | -34.2378 | -12.9286
7.3129 | -34.2382 | -12.9386
7.319 | -34.2385 | -12.9485
7.3251 | -34.2389 | -12.9583
7.3312 | -34.2392 | -12.9683
7.3373 | -34.2396 | -12.978
7.3434 | -34.24 | -12.9879
7.3495 | -34.2403 | -12.9976
7.3556 | -34.2407 | -13.0077
7.3617 | -34.2411 | -13.0178
7.3678 | -34.2414 | -13.0276
7.3739 | -34.2417 | -13.0377
7.3801 | -34.2421 | -13.0479
7.3862 | -34.2425 | -13.0577
7.3923 | -34.2428 | -13.0678
7.3984 | -34.2432 | -13.0776
7.4045 | -34.2435 | -13.0874
7.4106 | -34.2439 | -13.0973
7.4167 | -34.2442 | -13.1074
7.4228 | -34.2447 | -13.1174
7.4289 | -34.245 | -13.1273
7.435 | -34.2454 | -13.137
7.4411 | -34.2457 | -13.1469
7.4472 | -34.2461 | -13.1569
7.4533 | -34.2464 | -13.1666
7.4594 | -34.2468 | -13.1767
7.4655 | -34.2472 | -13.1865
7.4716 | -34.2475 | -13.1966
7.4777 | -34.2479 | -13.2065
7.4838 | -34.2482 | -13.2163
7.4899 | -34.2486 | -13.2261
7.496 | -34.2489 | -13.2358
7.5021 | -34.2493 | -13.2457
7.5082 | -34.2496 | -13.2556
7.5143 | -34.25 | -13.2656
7.5204 | -34.2503 | -13.2756
7.5266 | -34.2507 | -13.2853
7.5327 | -34.2511 | -13.2951
7.5388 | -34.2514 | -13.3053
7.5449 | -34.2518 | -13.3154
7.551 | -34.2522 | -13.3252
7.5571 | -34.2525 | -13.3352
7.5632 | -34.2529 | -13.3451
7.5693 | -34.2533 | -13.3549
7.5754 | -34.2536 | -13.3646
7.5815 | -34.254 | -13.3746
7.5876 | -34.2543 | -13.3847
7.5937 | -34.2547 | -13.3947
7.5998 | -34.2551 | -13.4045
7.6059 | -34.2555 | -13.4146
7.612 | -34.2558 | -13.4245
7.6181 | -34.2562 | -13.4343
7.6242 | -34.2566 | -13.4445
7.6303 | -34.2569 | -13.4545
7.6364 | -34.2573 | -13.4645
7.6425 | -34.2576 | -13.4745
7.6486 | -34.258 | -13.4844
7.6547 | -34.2584 | -13.4942
7.6608 | -34.2587 | -13.5041
7.667 | -34.2591 | -13.5141
7.6731 | -34.2594 | -13.5241
7.6792 | -34.2598 | -13.5342
7.6853 | -34.2602 | -13.5442
7.6914 | -34.2605 | -13.5543
7.6975 | -34.2609 | -13.5643
7.7036 | -34.2613 | -13.5742
7.7097 | -34.2616 | -13.5843
7.7158 | -34.262 | -13.5942
7.7219 | -34.2624 | -13.604
7.728 | -34.2628 | -13.614
7.7341 | -34.2631 | -13.624
7.7402 | -34.2635 | -13.634
7.7463 | -34.2639 | -13.6439
7.7524 | -34.2642 | -13.6541
7.7585 | -34.2646 | -13.664
7.7646 | -34.2649 | -13.674
7.7707 | -34.2653 | -13.684
7.7768 | -34.2656 | -13.6941
7.7829 | -34.266 | -13.7043
7.789 | -34.2664 | -13.7144
7.7951 | -34.2667 | -13.7244
7.8012 | -34.2671 | -13.7344
7.8073 | -34.2675 | -13.7445
7.8135 | -34.2679 | -13.7544
7.8196 | -34.2683 | -13.7646
7.8257 | -34.2686 | -13.7747
7.8318 | -34.269 | -13.7848
7.8379 | -34.2693 | -13.7946
7.844 | -34.2697 | -13.8045
7.8501 | -34.2701 | -13.8147
7.8562 | -34.2704 | -13.8249
7.8623 | -34.2708 | -13.8347
7.8684 | -34.2712 | -13.8447
7.8745 | -34.2715 | -13.8548
7.8806 | -34.2719 | -13.8649
7.8867 | -34.2722 | -13.8747
7.8928 | -34.2726 | -13.8845
7.8989 | -34.273 | -13.8945
7.905 | -34.2733 | -13.9045
7.9111 | -34.2737 | -13.9146
7.9172 | -34.2741 | -13.9245
7.9233 | -34.2744 | -13.9346
7.9294 | -34.2748 | -13.9447
7.9355 | -34.2751 | -13.9545
7.9416 | -34.2755 | -13.9645
7.9477 | -34.2758 | -13.9744
7.9539 | -34.2762 | -13.9843
7.96 | -34.2766 | -13.9943
7.9661 | -34.2769 | -14.0044
7.9722 | -34.2773 | -14.0142
7.9783 | -34.2777 | -14.0244
7.9844 | -34.2781 | -14.0343
7.9905 | -34.2784 | -14.0445
7.9966 | -34.2788 | -14.0546
8.0027 | -34.2791 | -14.0645
8.0088 | -34.2795 | -14.0745
8.0149 | -34.2798 | -14.0846
8.021 | -34.2802 | -14.0946
8.0271 | -34.2806 | -14.1048
8.0332 | -34.2809 | -14.1147
8.0393 | -34.2813 | -14.1248
8.0454 | -34.2817 | -14.135
8.0515 | -34.2821 | -14.145
8.0576 | -34.2824 | -14.1552
8.0637 | -34.2828 | -14.1651
8.0698 | -34.2831 | -14.175
8.0759 | -34.2835 | -14.1849
8.082 | -34.2839 | -14.1948
8.0881 | -34.2842 | -14.2049
8.0942 | -34.2846 | -14.215
8.1004 | -34.285 | -14.2251
8.1065 | -34.2853 | -14.235
8.1126 | -34.2857 | -14.2451
8.1187 | -34.286 | -14.2552
8.1248 | -34.2864 | -14.265
8.1309 | -34.2868 | -14.2749
8.137 | -34.2872 | -14.2847
8.1431 | -34.2875 | -14.2949
8.1492 | -34.2879 | -14.3049
8.1553 | -34.2882 | -14.3149
8.1614 | -34.2886 | -14.3251
8.1675 | -34.289 | -14.3351
8.1736 | -34.2893 | -14.3453
8.1797 | -34.2897 | -14.3553
8.1858 | -34.29 | -14.3654
8.1919 | -34.2904 | -14.3754
8.198 | -34.2908 | -14.3853
8.2041 | -34.2911 | -14.3954
8.2102 | -34.2915 | -14.4054
8.2163 | -34.2919 | -14.4153
8.2224 | -34.2922 | -14.4253
8.2285 | -34.2926 | -14.4351
8.2346 | -34.2929 | -14.445
8.2408 | -34.2933 | -14.4549
8.2469 | -34.2937 | -14.465
8.253 | -34.2941 | -14.4749
8.2591 | -34.2944 | -14.485
8.2652 | -34.2948 | -14.495
8.2713 | -34.2952 | -14.5051
8.2774 | -34.2956 | -14.5152
8.2835 | -34.2959 | -14.5254
8.2896 | -34.2963 | -14.5352
8.2957 | -34.2966 | -14.5454
8.3018 | -34.297 | -14.5554
8.3079 | -34.2974 | -14.5653
8.314 | -34.2978 | -14.5753
8.3201 | -34.2982 | -14.5852
8.3262 | -34.2985 | -14.5953
8.3323 | -34.2989 | -14.6052
8.3384 | -34.2993 | -14.6153
8.3445 | -34.2996 | -14.6252
8.3506 | -34.3 | -14.6353
8.3567 | -34.3004 | -14.6454
8.3628 | -34.3007 | -14.6554
8.3689 | -34.3011 | -14.6653
8.375 | -34.3014 | -14.6753
8.3812 | -34.3018 | -14.6852
8.3873 | -34.3022 | -14.6952
8.3934 | -34.3026 | -14.7052
8.3995 | -34.3029 | -14.7151
8.4056 | -34.3033 | -14.7251
8.4117 | -34.3036 | -14.7351
8.4178 | -34.304 | -14.7452
8.4239 | -34.3043 | -14.755
8.43 | -34.3047 | -14.7651
8.4361 | -34.305 | -14.7751
8.4422 | -34.3054 | -14.785
8.4483 | -34.3058 | -14.7949
8.4544 | -34.3062 | -14.8049
8.4605 | -34.3065 | -14.8148
8.4666 | -34.3069 | -14.8249
8.4727 | -34.3073 | -14.8349
8.4788 | -34.3076 | -14.8449
8.4849 | -34.308 | -14.8549
8.491 | -34.3083 | -14.8648
8.4971 | -34.3087 | -14.8747
8.5032 | -34.3091 | -14.8845
8.5093 | -34.3095 | -14.8943
8.5154 | -34.3098 | -14.9044
8.5215 | -34.3102 | -14.9142
8.5277 | -34.3106 | -14.9242
8.5338 | -34.311 | -14.9341
8.5399 | -34.3113 | -14.9441
8.546 | -34.3117 | -14.9539
8.5521 | -34.3121 | -14.9638
8.5582 | -34.3124 | -14.9738
8.5643 | -34.3128 | -14.9835
8.5704 | -34.3132 | -14.9933
8.5765 | -34.3135 | -15.0033
8.5826 | -34.3139 | -15.013
8.5887 | -34.3143 | -15.0229
8.5948 | -34.3147 | -15.0328
8.6009 | -34.3151 | -15.0426
8.607 | -34.3154 | -15.0526
8.6131 | -34.3158 | -15.0626
8.6192 | -34.3162 | -15.0724
8.6253 | -34.3166 | -15.0822
8.6314 | -34.3169 | -15.0921
8.6375 | -34.3173 | -15.1021
8.6436 | -34.3177 | -15.112
8.6497 | -34.318 | -15.1218
8.6558 | -34.3184 | -15.1317
8.6619 | -34.3187 | -15.1417
8.6681 | -34.3191 | -15.1518
8.6742 | -34.3195 | -15.1617
8.6803 | -34.3199 | -15.1715
8.6864 | -34.3203 | -15.1812
8.6925 | -34.3206 | -15.1909
8.6986 | -34.321 | -15.2009
8.7047 | -34.3214 | -15.2109
8.7108 | -34.3217 | -15.2208
8.7169 | -34.3221 | -15.2307
8.723 | -34.3225 | -15.2405
8.7291 | -34.3229 | -15.2502
8.7352 | -34.3233 | -15.2603
8.7413 | -34.3236 | -15.2702
8.7474 | -34.324 | -15.28
8.7535 | -34.3244 | -15.2898
8.7596 | -34.3248 | -15.2998
8.7657 | -34.3251 | -15.3098
8.7718 | -34.3255 | -15.3197
8.7779 | -34.3259 | -15.3297
8.784 | -34.3262 | -15.3396
8.7901 | -34.3267 | -15.3495
8.7962 | -34.327 | -15.3592
8.8023 | -34.3274 | -15.369
8.8084 | -34.3278 | -15.3788
8.8146 | -34.3282 | -15.3887
8.8207 | -34.3285 | -15.3986
8.8268 | -34.3289 | -15.4082
8.8329 | -34.3293 | -15.4181
8.839 | -34.3297 | -15.428
8.8451 | -34.33 | -15.4378
8.8512 | -34.3304 | -15.4475
8.8573 | -34.3308 | -15.4574
8.8634 | -34.3312 | -15.4671
8.8695 | -34.3316 | -15.4771
8.8756 | -34.3319 | -15.487
8.8817 | -34.3323 | -15.4968
8.8878 | -34.3327 | -15.5068
8.8939 | -34.3331 | -15.5166
8.9 | -34.3335 | -15.5264
8.9061 | -34.3339 | -15.5362
8.9122 | -34.3342 | -15.546
8.9183 | -34.3346 | -15.5559
8.9244 | -34.335 | -15.5659
8.9305 | -34.3355 | -15.5755
8.9366 | -34.3359 | -15.5853
8.9427 | -34.3362 | -15.5953
8.9488 | -34.3366 | -15.605
8.955 | -34.337 | -15.6148
8.9611 | -34.3374 | -15.6248
8.9672 | -34.3378 | -15.6346
8.9733 | -34.3382 | -15.6443
8.9794 | -34.3386 | -15.654
8.9855 | -34.339 | -15.664
8.9916 | -34.3394 | -15.6739
8.9977 | -34.3398 | -15.6837
9.0038 | -34.3402 | -15.6934
9.0099 | -34.3406 | -15.7032
9.016 | -34.3409 | -15.7131
9.0221 | -34.3413 | -15.7227
9.0282 | -34.3417 | -15.7323
9.0343 | -34.3421 | -15.7421
9.0404 | -34.3426 | -15.7519
9.0465 | -34.343 | -15.7619
9.0526 | -34.3434 | -15.7716
9.0587 | -34.3438 | -15.7814
9.0648 | -34.3442 | -15.7911
9.0709 | -34.3446 | -15.8008
9.077 | -34.345 | -15.8107
9.0831 | -34.3454 | -15.8206
9.0892 | -34.3458 | -15.8304
9.0953 | -34.3462 | -15.8403
9.1015 | -34.3466 | -15.8501
9.1076 | -34.347 | -15.8598
9.1137 | -34.3474 | -15.8695
9.1198 | -34.3478 | -15.8793
9.1259 | -34.3482 | -15.889
9.132 | -34.3486 | -15.899
9.1381 | -34.349 | -15.9086
9.1442 | -34.3495 | -15.9182
9.1503 | -34.3498 | -15.928
9.1564 | -34.3502 | -15.9378
9.1625 | -34.3506 | -15.9476
9.1686 | -34.3511 | -15.9574
9.1747 | -34.3515 | -15.9671
9.1808 | -34.3519 | -15.9769
9.1869 | -34.3523 | -15.9867
9.193 | -34.3527 | -15.9964
9.1991 | -34.3531 | -16.0063
9.2052 | -34.3535 | -16.016
9.2113 | -34.3539 | -16.0258
9.2174 | -34.3544 | -16.0356
9.2235 | -34.3548 | -16.0453
9.2296 | -34.3552 | -16.0551
9.2357 | -34.3556 | -16.065
9.2419 | -34.356 | -16.0747
9.248 | -34.3565 | -16.0844
9.2541 | -34.3569 | -16.0943
9.2602 | -34.3573 | -16.1041
9.2663 | -34.3577 | -16.1137
9.2724 | -34.3581 | -16.1234
9.2785 | -34.3585 | -16.1331
9.2846 | -34.359 | -16.1428
9.2907 | -34.3594 | -16.1525
9.2968 | -34.3598 | -16.1624
9.3029 | -34.3602 | -16.1723
9.309 | -34.3606 | -16.182
9.3151 | -34.361 | -16.1916
9.3212 | -34.3615 | -16.2012
9.3273 | -34.3619 | -16.2109
9.3334 | -34.3623 | -16.2206
9.3395 | -34.3627 | -16.2302
9.3456 | -34.3631 | -16.2398
9.3517 | -34.3635 | -16.2497
9.3578 | -34.364 | -16.2594
9.3639 | -34.3644 | -16.269
9.37 | -34.3648 | -16.2787
9.3761 | -34.3652 | -16.2886
9.3822 | -34.3657 | -16.2984
9.3884 | -34.3661 | -16.3079
9.3945 | -34.3665 | -16.3177
9.4006 | -34.367 | -16.3274
9.4067 | -34.3674 | -16.3372
9.4128 | -34.3678 | -16.3469
9.4189 | -34.3682 | -16.3568
9.425 | -34.3687 | -16.3665
9.4311 | -34.3691 | -16.3761
9.4372 | -34.3696 | -16.3858
9.4433 | -34.37 | -16.3955
9.4494 | -34.3704 | -16.4051
9.4555 | -34.3708 | -16.4149
9.4616 | -34.3713 | -16.4246
9.4677 | -34.3717 | -16.4343
9.4738 | -34.3721 | -16.444
9.4799 | -34.3726 | -16.4538
9.486 | -34.373 | -16.4635
9.4921 | -34.3734 | -16.4732
9.4982 | -34.3739 | -16.4827
9.5043 | -34.3743 | -16.4924
9.5104 | -34.3747 | -16.502
9.5165 | -34.3751 | -16.5117
9.5226 | -34.3755 | -16.5215
9.5288 | -34.376 | -16.531
9.5349 | -34.3764 | -16.5408
9.541 | -34.3768 | -16.5503
9.5471 | -34.3773 | -16.56
9.5532 | -34.3777 | -16.5698
9.5593 | -34.3781 | -16.5796
9.5654 | -34.3786 | -16.5892
9.5715 | -34.379 | -16.5989
9.5776 | -34.3794 | -16.6087
9.5837 | -34.3798 | -16.6183
9.5898 | -34.3802 | -16.6281
9.5959 | -34.3807 | -16.6379
9.602 | -34.3811 | -16.6473
9.6081 | -34.3815 | -16.657
9.6142 | -34.382 | -16.6666
9.6203 | -34.3824 | -16.6763
9.6264 | -34.3828 | -16.6859
9.6325 | -34.3832 | -16.6957
9.6386 | -34.3836 | -16.7053
9.6447 | -34.3841 | -16.715
9.6508 | -34.3845 | -16.7246
9.6569 | -34.3849 | -16.7343
9.663 | -34.3854 | -16.7441
9.6691 | -34.3858 | -16.7537
9.6753 | -34.3863 | -16.7633
9.6814 | -34.3867 | -16.7729
9.6875 | -34.3871 | -16.7824
9.6936 | -34.3875 | -16.7919
9.6997 | -34.388 | -16.8017
9.7058 | -34.3884 | -16.8112
9.7119 | -34.3889 | -16.821
9.718 | -34.3893 | -16.8308
9.7241 | -34.3897 | -16.8403
9.7302 | -34.3902 | -16.85
9.7363 | -34.3906 | -16.8596
9.7424 | -34.3911 | -16.8691
9.7485 | -34.3915 | -16.8788
9.7546 | -34.3919 | -16.8883
9.7607 | -34.3923 | -16.8981
9.7668 | -34.3928 | -16.9078
9.7729 | -34.3932 | -16.9174
9.779 | -34.3937 | -16.9268
9.7851 | -34.3941 | -16.9366
9.7912 | -34.3946 | -16.9462
9.7973 | -34.395 | -16.9559
9.8034 | -34.3955 | -16.9653
9.8095 | -34.3959 | -16.9751
9.8157 | -34.3964 | -16.9847
9.8218 | -34.3968 | -16.9946
9.8279 | -34.3972 | -17.0041
9.834 | -34.3976 | -17.0138
9.8401 | -34.3981 | -17.0235
9.8462 | -34.3985 | -17.0329
9.8523 | -34.3989 | -17.0426
9.8584 | -34.3994 | -17.0521
9.8645 | -34.3998 | -17.0619
9.8706 | -34.4002 | -17.0716
9.8767 | -34.4006 | -17.0812
9.8828 | -34.4011 | -17.0906
9.8889 | -34.4015 | -17.1
9.895 | -34.4019 | -17.1097
9.9011 | -34.4024 | -17.1193
9.9072 | -34.4028 | -17.1289
9.9133 | -34.4032 | -17.1385
9.9194 | -34.4037 | -17.148
9.9255 | -34.4041 | -17.1576
9.9316 | -34.4045 | -17.1671
9.9377 | -34.405 | -17.1767
9.9438 | -34.4054 | -17.1863
9.9499 | -34.4058 | -17.1961
9.956 | -34.4063 | -17.2057
9.9622 | -34.4067 | -17.2154
9.9683 | -34.4072 | -17.225
9.9744 | -34.4076 | -17.2347
9.9805 | -34.408 | -17.2443
9.9866 | -34.4084 | -17.2538
9.9927 | -34.4089 | -17.2633
9.9988 | -34.4093 | -17.2731
10.0049 | -34.4098 | -17.2826
10.011 | -34.4102 | -17.2923
10.0171 | -34.4106 | -17.3018
10.0232 | -34.4111 | -17.3113
10.0293 | -34.4115 | -17.3209
10.0354 | -34.412 | -17.3305
10.0415 | -34.4124 | -17.3401
10.0476 | -34.4129 | -17.3499
10.0537 | -34.4133 | -17.3595
10.0598 | -34.4137 | -17.369
10.0659 | -34.4142 | -17.3786
10.072 | -34.4146 | -17.388
10.0781 | -34.415 | -17.3975
10.0842 | -34.4155 | -17.4071
10.0903 | -34.4159 | -17.4166
10.0964 | -34.4164 | -17.426
10.1026 | -34.4168 | -17.4356
10.1087 | -34.4172 | -17.4452
10.1148 | -34.4176 | -17.4548
10.1209 | -34.4181 | -17.4642
10.127 | -34.4185 | -17.4736
10.1331 | -34.419 | -17.483
10.1392 | -34.4194 | -17.4925
10.1453 | -34.4198 | -17.502
10.1514 | -34.4203 | -17.5115
10.1575 | -34.4207 | -17.5212
10.1636 | -34.4211 | -17.5305
10.1697 | -34.4216 | -17.5402
10.1758 | -34.422 | -17.5495
10.1819 | -34.4225 | -17.559
10.188 | -34.4229 | -17.5687
10.1941 | -34.4233 | -17.5782
10.2002 | -34.4237 | -17.5877
10.2063 | -34.4241 | -17.5972
10.2124 | -34.4246 | -17.6068
10.2185 | -34.425 | -17.6164
10.2246 | -34.4255 | -17.6258
10.2307 | -34.4259 | -17.6354
10.2368 | -34.4263 | -17.6446
10.2429 | -34.4267 | -17.6541
10.2491 | -34.4271 | -17.6637
10.2552 | -34.4276 | -17.6733
10.2613 | -34.428 | -17.6829
10.2674 | -34.4285 | -17.6925
10.2735 | -34.4289 | -17.7022
10.2796 | -34.4294 | -17.7116
10.2857 | -34.4298 | -17.7212
10.2918 | -34.4302 | -17.7308
10.2979 | -34.4307 | -17.7403
10.304 | -34.4311 | -17.7496
10.3101 | -34.4315 | -17.7592
10.3162 | -34.432 | -17.7688
10.3223 | -34.4324 | -17.7782
10.3284 | -34.4329 | -17.7877
10.3345 | -34.4333 | -17.7971
10.3406 | -34.4337 | -17.8065
10.3467 | -34.4341 | -17.816
10.3528 | -34.4346 | -17.8253
10.3589 | -34.435 | -17.8348
10.365 | -34.4354 | -17.8445
10.3711 | -34.4359 | -17.8541
10.3772 | -34.4363 | -17.8634
10.3833 | -34.4368 | -17.8729
10.3895 | -34.4372 | -17.8824
10.3956 | -34.4376 | -17.8919
10.4017 | -34.4381 | -17.9013
10.4078 | -34.4385 | -17.9107
10.4139 | -34.4389 | -17.9203
10.42 | -34.4394 | -17.9299
10.4261 | -34.4398 | -17.9393
10.4322 | -34.4402 | -17.9489
10.4383 | -34.4407 | -17.9583
10.4444 | -34.4411 | -17.9675
10.4505 | -34.4415 | -17.9772
10.4566 | -34.4419 | -17.9866
10.4627 | -34.4423 | -17.996
10.4688 | -34.4428 | -18.0054
10.4749 | -34.4432 | -18.0149
10.481 | -34.4436 | -18.0244
10.4871 | -34.444 | -18.0337
10.4932 | -34.4445 | -18.043
10.4993 | -34.4449 | -18.0525
10.5054 | -34.4453 | -18.062
10.5115 | -34.4458 | -18.0714
10.5176 | -34.4462 | -18.0806
10.5237 | -34.4467 | -18.0901
10.5298 | -34.4471 | -18.0994
10.536 | -34.4475 | -18.1088
10.5421 | -34.4479 | -18.1182
10.5482 | -34.4483 | -18.1277
10.5543 | -34.4488 | -18.1369
10.5604 | -34.4492 | -18.1462
10.5665 | -34.4497 | -18.1555
10.5726 | -34.4501 | -18.165
10.5787 | -34.4505 | -18.1745
10.5848 | -34.451 | -18.1836
10.5909 | -34.4514 | -18.1932
10.597 | -34.4518 | -18.2026
10.6031 | -34.4523 | -18.2122
10.6092 | -34.4527 | -18.2216
10.6153 | -34.4531 | -18.2312
10.6214 | -34.4536 | -18.2404
10.6275 | -34.454 | -18.2499
10.6336 | -34.4544 | -18.2593
10.6397 | -34.4548 | -18.2689
10.6458 | -34.4553 | -18.2783
10.6519 | -34.4557 | -18.2877
10.658 | -34.4561 | -18.297
10.6641 | -34.4566 | -18.3065
10.6702 | -34.457 | -18.316
10.6764 | -34.4575 | -18.3253
10.6825 | -34.4579 | -18.3346
10.6886 | -34.4583 | -18.344
10.6947 | -34.4587 | -18.3533
10.7008 | -34.4592 | -18.3627
10.7069 | -34.4596 | -18.3722
10.713 | -34.4601 | -18.3815
10.7191 | -34.4605 | -18.3908
10.7252 | -34.461 | -18.3999
10.7313 | -34.4614 | -18.4092
10.7374 | -34.4618 | -18.4187
10.7435 | -34.4623 | -18.4282
10.7496 | -34.4626 | -18.4373
10.7557 | -34.4631 | -18.4468
10.7618 | -34.4636 | -18.4558
10.7679 | -34.464 | -18.4651
10.774 | -34.4644 | -18.4746
10.7801 | -34.4648 | -18.4839
10.7862 | -34.4652 | -18.4931
10.7923 | -34.4657 | -18.5025
10.7984 | -34.4661 | -18.5117
10.8045 | -34.4666 | -18.521
10.8106 | -34.4669 | -18.5304
10.8168 | -34.4674 | -18.5394
10.8229 | -34.4678 | -18.5489
10.829 | -34.4682 | -18.5583
10.8351 | -34.4687 | -18.5675
10.8412 | -34.4692 | -18.5767
10.8473 | -34.4697 | -18.5857
10.8534 | -34.4701 | -18.5948
10.8595 | -34.4706 | -18.6043
10.8656 | -34.4711 | -18.6135
10.8717 | -34.4715 | -18.6228
10.8778 | -34.4719 | -18.632
10.8839 | -34.4724 | -18.6413
10.89 | -34.4728 | -18.6507
10.8961 | -34.4733 | -18.6601
10.9022 | -34.4737 | -18.6693
10.9083 | -34.4742 | -18.6784
10.9144 | -34.4746 | -18.6875
10.9205 | -34.4751 | -18.6968
10.9266 | -34.4755 | -18.7061
10.9327 | -34.476 | -18.7153
10.9388 | -34.4764 | -18.7243
10.9449 | -34.4769 | -18.7337
10.951 | -34.4773 | -18.7429
10.9571 | -34.4777 | -18.7521
10.9633 | -34.4782 | -18.7614
10.9694 | -34.4786 | -18.7707
10.9755 | -34.4791 | -18.7797
10.9816 | -34.4795 | -18.7889
10.9877 | -34.48 | -18.7979
10.9938 | -34.4804 | -18.8072
10.9999 | -34.4809 | -18.8162
11.006 | -34.4813 | -18.8252
11.0121 | -34.4818 | -18.8344
11.0182 | -34.4822 | -18.8435
11.0243 | -34.4826 | -18.8528
11.0304 | -34.483 | -18.8621
11.0365 | -34.4835 | -18.871
11.0426 | -34.4839 | -18.8804
11.0487 | -34.4844 | -18.89
11.0548 | -34.4849 | -18.899
11.0609 | -34.4853 | -18.9085
11.067 | -34.4858 | -18.9179
11.0731 | -34.4862 | -18.9274
11.0792 | -34.4867 | -18.9365
11.0853 | -34.4871 | -18.9456
11.0914 | -34.4876 | -18.9552
11.0975 | -34.488 | -18.9643
11.1037 | -34.4884 | -18.9736
11.1098 | -34.4888 | -18.9828
11.1159 | -34.4893 | -18.9916
11.122 | -34.4897 | -19.001
11.1281 | -34.4901 | -19.0107
11.1342 | -34.4906 | -19.0199
11.1403 | -34.4911 | -19.0292
11.1464 | -34.4915 | -19.0382
11.1525 | -34.492 | -19.0469
11.1586 | -34.4924 | -19.0566
11.1647 | -34.4928 | -19.0659
11.1708 | -34.4933 | -19.075
11.1769 | -34.4938 | -19.0842
11.183 | -34.4943 | -19.0937
11.1891 | -34.4947 | -19.1032
11.1952 | -34.4952 | -19.1125
11.2013 | -34.4956 | -19.1215
11.2074 | -34.496 | -19.1308
11.2135 | -34.4965 | -19.1401
11.2196 | -34.497 | -19.1495
11.2257 | -34.4975 | -19.1588
11.2318 | -34.498 | -19.1679
11.2379 | -34.4984 | -19.1773
11.244 | -34.4989 | -19.1865
11.2502 | -34.4994 | -19.1956
11.2563 | -34.4998 | -19.2044
11.2624 | -34.5003 | -19.2135
11.2685 | -34.5008 | -19.2227
11.2746 | -34.5013 | -19.2315
11.2807 | -34.5017 | -19.2409
11.2868 | -34.5022 | -19.2502
11.2929 | -34.5027 | -19.2593
11.299 | -34.5032 | -19.2689
11.3051 | -34.5036 | -19.2789
11.3112 | -34.5041 | -19.2886
11.3173 | -34.5045 | -19.2973
11.3234 | -34.5049 | -19.3066
11.3295 | -34.5053 | -19.3161
11.3356 | -34.5057 | -19.3253
11.3417 | -34.5062 | -19.3345
11.3478 | -34.5066 | -19.3436
11.3539 | -34.507 | -19.3523
11.36 | -34.5075 | -19.3612
11.3661 | -34.5079 | -19.37
11.3722 | -34.5084 | -19.3788
11.3783 | -34.5089 | -19.3881
11.3844 | -34.5093 | -19.3972
11.3906 | -34.5098 | -19.4063
11.3967 | -34.5103 | -19.4152
11.4028 | -34.5107 | -19.4244
11.4089 | -34.5111 | -19.4334
11.415 | -34.5116 | -19.4426
11.4211 | -34.5121 | -19.4519
11.4272 | -34.5125 | -19.4613
11.4333 | -34.513 | -19.4707
11.4394 | -34.5134 | -19.4798
11.4455 | -34.5139 | -19.489
11.4516 | -34.5143 | -19.4984
11.4577 | -34.5148 | -19.5075
11.4638 | -34.5153 | -19.5161
11.4699 | -34.5158 | -19.5253
11.476 | -34.5163 | -19.5345
11.4821 | -34.5167 | -19.5437
11.4882 | -34.5172 | -19.5528
11.4943 | -34.5176 | -19.562
11.5004 | -34.5181 | -19.5708
11.5065 | -34.5186 | -19.58
11.5126 | -34.519 | -19.5887
11.5187 | -34.5195 | -19.5978
11.5248 | -34.52 | -19.6069
11.5309 | -34.5205 | -19.616
11.5371 | -34.5209 | -19.6254
11.5432 | -34.5214 | -19.6346
11.5493 | -34.5219 | -19.6437
11.5554 | -34.5223 | -19.6527
11.5615 | -34.5228 | -19.6618
11.5676 | -34.5232 | -19.671
11.5737 | -34.5237 | -19.6802
11.5798 | -34.5242 | -19.6892
11.5859 | -34.5247 | -19.6983
11.592 | -34.5252 | -19.7074
11.5981 | -34.5256 | -19.7166
11.6042 | -34.5261 | -19.7256
11.6103 | -34.5265 | -19.7346
11.6164 | -34.527 | -19.7436
11.6225 | -34.5275 | -19.7527
11.6286 | -34.5279 | -19.762
11.6347 | -34.5284 | -19.7711
11.6408 | -34.5289 | -19.7803
11.6469 | -34.5294 | -19.7895
11.653 | -34.5298 | -19.7984
11.6591 | -34.5303 | -19.8073
11.6652 | -34.5307 | -19.8163
11.6713 | -34.5312 | -19.8252
11.6775 | -34.5317 | -19.8339
11.6836 | -34.5321 | -19.843
11.6897 | -34.5326 | -19.852
11.6958 | -34.5331 | -19.8612
11.7019 | -34.5335 | -19.8703
11.708 | -34.534 | -19.8792
11.7141 | -34.5345 | -19.8887
11.7202 | -34.5349 | -19.8977
11.7263 | -34.5354 | -19.907
11.7324 | -34.5359 | -19.9163
11.7385 | -34.5363 | -19.9253
11.7446 | -34.5367 | -19.9343
11.7507 | -34.5372 | -19.9433
11.7568 | -34.5376 | -19.9524
11.7629 | -34.5381 | -19.9614
11.769 | -34.5386 | -19.9704
11.7751 | -34.539 | -19.9797
11.7812 | -34.5396 | -19.9885
11.7873 | -34.54 | -19.9978
11.7934 | -34.5405 | -20.0068
11.7995 | -34.541 | -20.0158
11.8056 | -34.5414 | -20.0248
11.8117 | -34.5419 | -20.034
11.8178 | -34.5424 | -20.0432
11.824 | -34.5429 | -20.0522
11.8301 | -34.5433 | -20.0612
11.8362 | -34.5438 | -20.0702
11.8423 | -34.5443 | -20.0792
11.8484 | -34.5447 | -20.0882
11.8545 | -34.5452 | -20.097
11.8606 | -34.5457 | -20.1059
11.8667 | -34.5462 | -20.1149
11.8728 | -34.5467 | -20.1239
11.8789 | -34.5471 | -20.1329
11.885 | -34.5476 | -20.1421
11.8911 | -34.5481 | -20.1509
11.8972 | -34.5486 | -20.1599
11.9033 | -34.5491 | -20.1689
11.9094 | -34.5495 | -20.1778
11.9155 | -34.55 | -20.1868
11.9216 | -34.5505 | -20.1958
11.9277 | -34.5509 | -20.2047
11.9338 | -34.5514 | -20.2138
11.9399 | -34.5519 | -20.2228
11.946 | -34.5523 | -20.2318
11.9521 | -34.5528 | -20.2407
11.9582 | -34.5533 | -20.2497
11.9644 | -34.5537 | -20.2587
11.9705 | -34.5542 | -20.2677
11.9766 | -34.5547 | -20.2767
11.9827 | -34.5552 | -20.2858
11.9888 | -34.5557 | -20.2947
11.9949 | -34.5562 | -20.3036
12.001 | -34.5566 | -20.3125
12.0071 | -34.5571 | -20.3215
12.0132 | -34.5576 | -20.3303
12.0193 | -34.558 | -20.3394
12.0254 | -34.5585 | -20.3483
12.0315 | -34.559 | -20.3571
12.0376 | -34.5595 | -20.3658
12.0437 | -34.5599 | -20.3748
12.0498 | -34.5604 | -20.3839
12.0559 | -34.5609 | -20.3929
12.062 | -34.5614 | -20.4018
12.0681 | -34.5619 | -20.4108
12.0742 | -34.5624 | -20.4195
12.0803 | -34.5629 | -20.4287
12.0864 | -34.5633 | -20.4376
12.0925 | -34.5638 | -20.4465
12.0986 | -34.5643 | -20.4552
12.1047 | -34.5647 | -20.4642
12.1109 | -34.5652 | -20.473
12.117 | -34.5657 | -20.4818
12.1231 | -34.5662 | -20.4906
12.1292 | -34.5666 | -20.4996
12.1353 | -34.5671 | -20.5086
12.1414 | -34.5676 | -20.5178
12.1475 | -34.5681 | -20.5267
12.1536 | -34.5685 | -20.5357
12.1597 | -34.569 | -20.5448
12.1658 | -34.5695 | -20.5536
12.1719 | -34.57 | -20.5624
12.178 | -34.5705 | -20.5713
12.1841 | -34.5709 | -20.5804
12.1902 | -34.5714 | -20.5893
12.1963 | -34.5719 | -20.5984
12.2024 | -34.5724 | -20.6069
12.2085 | -34.5729 | -20.616
12.2146 | -34.5734 | -20.6249
12.2207 | -34.5739 | -20.6338
12.2268 | -34.5743 | -20.6427
12.2329 | -34.5748 | -20.6516
12.239 | -34.5753 | -20.6602
12.2451 | -34.5758 | -20.6691
12.2513 | -34.5762 | -20.678
12.2574 | -34.5767 | -20.687
12.2635 | -34.5772 | -20.6958
12.2696 | -34.5777 | -20.7048
12.2757 | -34.5781 | -20.7133
12.2818 | -34.5786 | -20.7222
12.2879 | -34.5791 | -20.731
12.294 | -34.5795 | -20.7399
12.3001 | -34.58 | -20.7485
12.3062 | -34.5805 | -20.7574
12.3123 | -34.5809 | -20.766
12.3184 | -34.5814 | -20.7747
12.3245 | -34.5819 | -20.7834
12.3306 | -34.5823 | -20.7922
12.3367 | -34.5828 | -20.801
12.3428 | -34.5833 | -20.8099
12.3489 | -34.5838 | -20.8183
12.355 | -34.5843 | -20.827
12.3611 | -34.5848 | -20.836
12.3672 | -34.5853 | -20.8447
12.3733 | -34.5857 | -20.8537
12.3794 | -34.5863 | -20.8624
12.3855 | -34.5867 | -20.8711
12.3916 | -34.5872 | -20.8798
12.3978 | -34.5877 | -20.8887
12.4039 | -34.5882 | -20.8976
12.41 | -34.5886 | -20.9064
12.4161 | -34.5891 | -20.9148
12.4222 | -34.5896 | -20.9236
12.4283 | -34.5901 | -20.9323
12.4344 | -34.5906 | -20.9412
12.4405 | -34.591 | -20.9501
12.4466 | -34.5915 | -20.9588
12.4527 | -34.592 | -20.9675
12.4588 | -34.5925 | -20.9763
12.4649 | -34.5929 | -20.9851
12.471 | -34.5934 | -20.9939
12.4771 | -34.5939 | -21.0026
12.4832 | -34.5943 | -21.0112
12.4893 | -34.5948 | -21.0198
12.4954 | -34.5953 | -21.0284
12.5015 | -34.5957 | -21.0371
12.5076 | -34.5962 | -21.0459
12.5137 | -34.5967 | -21.0547
12.5198 | -34.5972 | -21.0633
12.5259 | -34.5977 | -21.0722
12.532 | -34.5981 | -21.081
12.5382 | -34.5986 | -21.0898
12.5443 | -34.599 | -21.0985
12.5504 | -34.5995 | -21.1073
12.5565 | -34.6 | -21.1158
12.5626 | -34.6004 | -21.1246
12.5687 | -34.6009 | -21.1334
12.5748 | -34.6014 | -21.1422
12.5809 | -34.6019 | -21.151
12.587 | -34.6023 | -21.1598
12.5931 | -34.6029 | -21.1684
12.5992 | -34.6033 | -21.1772
12.6053 | -34.6038 | -21.1859
12.6114 | -34.6043 | -21.1946
12.6175 | -34.6048 | -21.2034
12.6236 | -34.6052 | -21.212
12.6297 | -34.6057 | -21.2206
12.6358 | -34.6062 | -21.2291
12.6419 | -34.6066 | -21.2378
12.648 | -34.6071 | -21.2465
12.6541 | -34.6076 | -21.2552
12.6602 | -34.6081 | -21.2636
12.6663 | -34.6086 | -21.2723
12.6724 | -34.6091 | -21.2809
12.6785 | -34.6095 | -21.2895
12.6847 | -34.61 | -21.2981
12.6908 | -34.6104 | -21.3067
12.6969 | -34.6109 | -21.3151
12.703 | -34.6114 | -21.3238
12.7091 | -34.6118 | -21.3324
12.7152 | -34.6123 | -21.3409
12.7213 | -34.6128 | -21.3495
12.7274 | -34.6133 | -21.3584
12.7335 | -34.6137 | -21.3667
12.7396 | -34.6142 | -21.3754
12.7457 | -34.6146 | -21.3839
12.7518 | -34.6151 | -21.3925
12.7579 | -34.6156 | -21.4013
12.764 | -34.616 | -21.4101
12.7701 | -34.6165 | -21.4187
12.7762 | -34.617 | -21.4272
12.7823 | -34.6174 | -21.4358
12.7884 | -34.6179 | -21.4445
12.7945 | -34.6184 | -21.4531
12.8006 | -34.6189 | -21.4617
12.8067 | -34.6194 | -21.4704
12.8128 | -34.6199 | -21.4789
12.8189 | -34.6204 | -21.4874
12.8251 | -34.6208 | -21.4962
12.8312 | -34.6213 | -21.5048
12.8373 | -34.6217 | -21.5132
12.8434 | -34.6222 | -21.5219
12.8495 | -34.6226 | -21.5305
12.8556 | -34.6231 | -21.5391
12.8617 | -34.6236 | -21.5479
12.8678 | -34.6241 | -21.5567
12.8739 | -34.6245 | -21.5649
12.88 | -34.625 | -21.5734
12.8861 | -34.6254 | -21.5819
12.8922 | -34.6259 | -21.5906
12.8983 | -34.6264 | -21.5992
12.9044 | -34.6269 | -21.6078
12.9105 | -34.6273 | -21.6162
12.9166 | -34.6278 | -21.6248
12.9227 | -34.6282 | -21.6334
12.9288 | -34.6287 | -21.642
12.9349 | -34.6292 | -21.6507
12.941 | -34.6297 | -21.6591
12.9471 | -34.6301 | -21.6678
12.9532 | -34.6306 | -21.6763
12.9593 | -34.6311 | -21.685
12.9654 | -34.6315 | -21.6937
12.9716 | -34.632 | -21.7024
12.9777 | -34.6325 | -21.7109
12.9838 | -34.633 | -21.7193
12.9899 | -34.6334 | -21.7277
12.996 | -34.6339 | -21.7362
13.0021 | -34.6344 | -21.7446
13.0082 | -34.6349 | -21.753
13.0143 | -34.6353 | -21.7613
13.0204 | -34.6358 | -21.7698
13.0265 | -34.6363 | -21.7785
13.0326 | -34.6367 | -21.7871
13.0387 | -34.6372 | -21.7958
13.0448 | -34.6376 | -21.8045
13.0509 | -34.6381 | -21.8126
13.057 | -34.6386 | -21.821
13.0631 | -34.6391 | -21.8295
13.0692 | -34.6396 | -21.8381
13.0753 | -34.64 | -21.8467
13.0814 | -34.6405 | -21.8554
13.0875 | -34.641 | -21.8639
13.0936 | -34.6414 | -21.8724
13.0997 | -34.6419 | -21.8809
13.1058 | -34.6424 | -21.8893
13.112 | -34.6429 | -21.8978
13.1181 | -34.6433 | -21.9062
13.1242 | -34.6438 | -21.9146
13.1303 | -34.6443 | -21.9233
13.1364 | -34.6447 | -21.9318
13.1425 | -34.6451 | -21.9403
13.1486 | -34.6456 | -21.949
13.1547 | -34.6461 | -21.9573
13.1608 | -34.6466 | -21.9656
13.1669 | -34.6471 | -21.974
13.173 | -34.6476 | -21.9825
13.1791 | -34.648 | -21.9907
13.1852 | -34.6485 | -21.9993
13.1913 | -34.6489 | -22.0077
13.1974 | -34.6494 | -22.0162
13.2035 | -34.6498 | -22.0245
13.2096 | -34.6503 | -22.0329
13.2157 | -34.6508 | -22.0412
13.2218 | -34.6512 | -22.0497
13.2279 | -34.6517 | -22.0578
13.234 | -34.6522 | -22.0662
13.2401 | -34.6527 | -22.0747
13.2462 | -34.6531 | -22.0831
13.2524 | -34.6536 | -22.0917
13.2585 | -34.6541 | -22.1
13.2646 | -34.6545 | -22.1085
13.2707 | -34.655 | -22.1168
13.2768 | -34.6555 | -22.1254
13.2829 | -34.656 | -22.1342
13.289 | -34.6564 | -22.1425
13.2951 | -34.6569 | -22.151
13.3012 | -34.6573 | -22.1594
13.3073 | -34.6578 | -22.168
13.3134 | -34.6582 | -22.1768
13.3195 | -34.6587 | -22.1851
13.3256 | -34.6591 | -22.1936
13.3317 | -34.6596 | -22.202
13.3378 | -34.6601 | -22.2104
13.3439 | -34.6606 | -22.219
13.35 | -34.661 | -22.2277
13.3561 | -34.6615 | -22.236
13.3622 | -34.662 | -22.2444
13.3683 | -34.6624 | -22.2526
13.3744 | -34.6629 | -22.2609
13.3805 | -34.6634 | -22.2694
13.3866 | -34.6638 | -22.2777
13.3927 | -34.6642 | -22.2863
13.3989 | -34.6647 | -22.2946
13.405 | -34.6652 | -22.303
13.4111 | -34.6656 | -22.3113
13.4172 | -34.6661 | -22.32
13.4233 | -34.6665 | -22.3287
13.4294 | -34.667 | -22.3372
13.4355 | -34.6675 | -22.3455
13.4416 | -34.668 | -22.3539
13.4477 | -34.6684 | -22.3624
13.4538 | -34.6689 | -22.3707
13.4599 | -34.6694 | -22.3792
13.466 | -34.6698 | -22.3876
13.4721 | -34.6702 | -22.3963
13.4782 | -34.6707 | -22.4046
13.4843 | -34.6712 | -22.4131
13.4904 | -34.6716 | -22.4215
13.4965 | -34.6721 | -22.4301
13.5026 | -34.6725 | -22.4386
13.5087 | -34.673 | -22.4469
13.5148 | -34.6735 | -22.4553
13.5209 | -34.6739 | -22.464
13.527 | -34.6744 | -22.4724
13.5331 | -34.6749 | -22.4808
13.5393 | -34.6753 | -22.4893
13.5454 | -34.6757 | -22.4977
13.5515 | -34.6762 | -22.5061
13.5576 | -34.6766 | -22.5143
13.5637 | -34.6771 | -22.5228
13.5698 | -34.6776 | -22.5313
13.5759 | -34.678 | -22.5396
13.582 | -34.6785 | -22.5479
13.5881 | -34.679 | -22.5563
13.5942 | -34.6794 | -22.5649
13.6003 | -34.6799 | -22.5733
13.6064 | -34.6804 | -22.5818
13.6125 | -34.6809 | -22.5902
13.6186 | -34.6813 | -22.5986
13.6247 | -34.6818 | -22.6072
13.6308 | -34.6823 | -22.6157
13.6369 | -34.6827 | -22.6243
13.643 | -34.6832 | -22.6329
13.6491 | -34.6837 | -22.6412
13.6552 | -34.6842 | -22.6497
13.6613 | -34.6846 | -22.6579
13.6674 | -34.6851 | -22.6664
13.6735 | -34.6855 | -22.6751
13.6796 | -34.686 | -22.6836
13.6858 | -34.6865 | -22.6919
13.6919 | -34.6869 | -22.7003
13.698 | -34.6874 | -22.7088
13.7041 | -34.6879 | -22.7176
13.7102 | -34.6883 | -22.7259
13.7163 | -34.6888 | -22.734
13.7224 | -34.6892 | -22.7425
13.7285 | -34.6897 | -22.7511
13.7346 | -34.6902 | -22.7597
13.7407 | -34.6907 | -22.7679
13.7468 | -34.6911 | -22.7762
13.7529 | -34.6916 | -22.7846
13.759 | -34.6921 | -22.7932
13.7651 | -34.6925 | -22.8016
13.7712 | -34.693 | -22.8099
13.7773 | -34.6935 | -22.8184
13.7834 | -34.6939 | -22.8269
13.7895 | -34.6944 | -22.8352
13.7956 | -34.6949 | -22.8437
13.8017 | -34.6953 | -22.8521
13.8078 | -34.6958 | -22.8606
13.8139 | -34.6962 | -22.8688
13.82 | -34.6967 | -22.8772
13.8262 | -34.6972 | -22.8857
13.8323 | -34.6976 | -22.8941
13.8384 | -34.6981 | -22.9026
13.8445 | -34.6986 | -22.911
13.8506 | -34.6991 | -22.9194
13.8567 | -34.6996 | -22.9277
13.8628 | -34.7001 | -22.9361
13.8689 | -34.7006 | -22.9446
13.875 | -34.701 | -22.953
13.8811 | -34.7014 | -22.9614
13.8872 | -34.7019 | -22.9698
13.8933 | -34.7023 | -22.9783
13.8994 | -34.7028 | -22.9868
13.9055 | -34.7032 | -22.9955
13.9116 | -34.7037 | -23.0039
13.9177 | -34.7042 | -23.0124
13.9238 | -34.7047 | -23.0209
13.9299 | -34.7052 | -23.0294
13.936 | -34.7056 | -23.0379
13.9421 | -34.7061 | -23.0463
13.9482 | -34.7065 | -23.0547
13.9543 | -34.707 | -23.0633
13.9604 | -34.7074 | -23.0717
13.9665 | -34.7079 | -23.0801
13.9727 | -34.7084 | -23.0886
13.9788 | -34.7089 | -23.0969
13.9849 | -34.7094 | -23.1053
13.991 | -34.7098 | -23.1138
13.9971 | -34.7103 | -23.1222
14.0032 | -34.7107 | -23.1306
14.0093 | -34.7112 | -23.1391
14.0154 | -34.7117 | -23.1475
14.0215 | -34.7122 | -23.1557
14.0276 | -34.7126 | -23.164
14.0337 | -34.7131 | -23.1724
14.0398 | -34.7136 | -23.1807
14.0459 | -34.7141 | -23.1891
14.052 | -34.7145 | -23.1976
14.0581 | -34.715 | -23.2061
14.0642 | -34.7154 | -23.2146
14.0703 | -34.7159 | -23.223
14.0764 | -34.7164 | -23.2314
14.0825 | -34.7169 | -23.2399
14.0886 | -34.7174 | -23.2484
14.0947 | -34.7178 | -23.2568
14.1008 | -34.7183 | -23.2652
14.1069 | -34.7188 | -23.2737
14.1131 | -34.7193 | -23.2823
14.1192 | -34.7197 | -23.2906
14.1253 | -34.7202 | -23.2989
14.1314 | -34.7207 | -23.3074
14.1375 | -34.7212 | -23.3158
14.1436 | -34.7216 | -23.3242
14.1497 | -34.7221 | -23.3326
14.1558 | -34.7226 | -23.3409
14.1619 | -34.7231 | -23.3493
14.168 | -34.7235 | -23.3578
14.1741 | -34.724 | -23.3662
14.1802 | -34.7245 | -23.3746
14.1863 | -34.725 | -23.3831
14.1924 | -34.7255 | -23.3916
14.1985 | -34.7259 | -23.4001
14.2046 | -34.7264 | -23.4086
14.2107 | -34.7269 | -23.4169
14.2168 | -34.7274 | -23.4254
14.2229 | -34.7278 | -23.4339
14.229 | -34.7283 | -23.4424
14.2351 | -34.7288 | -23.4509
14.2412 | -34.7293 | -23.4593
14.2473 | -34.7298 | -23.4678
14.2534 | -34.7302 | -23.4762
14.2596 | -34.7307 | -23.4846
14.2657 | -34.7312 | -23.493
14.2718 | -34.7316 | -23.5014
14.2779 | -34.7321 | -23.51
14.284 | -34.7326 | -23.5185
14.2901 | -34.7331 | -23.5269
14.2962 | -34.7336 | -23.5355
14.3023 | -34.734 | -23.544
14.3084 | -34.7345 | -23.5524
14.3145 | -34.735 | -23.561
14.3206 | -34.7355 | -23.5695
14.3267 | -34.736 | -23.578
14.3328 | -34.7364 | -23.5865
14.3389 | -34.7369 | -23.5948
14.345 | -34.7374 | -23.6033
14.3511 | -34.7379 | -23.6119
14.3572 | -34.7383 | -23.6204
14.3633 | -34.7388 | -23.6289
14.3694 | -34.7393 | -23.6373
14.3755 | -34.7398 | -23.6458
14.3816 | -34.7403 | -23.6543
14.3877 | -34.7408 | -23.6629
14.3938 | -34.7413 | -23.6716
14.4 | -34.7418 | -23.68
14.4061 | -34.7422 | -23.6884
14.4122 | -34.7427 | -23.6968
14.4183 | -34.7432 | -23.7052
14.4244 | -34.7437 | -23.7137
14.4305 | -34.7442 | -23.7221
14.4366 | -34.7447 | -23.7306
14.4427 | -34.7452 | -23.7392
14.4488 | -34.7456 | -23.7477
14.4549 | -34.7461 | -23.7561
14.461 | -34.7466 | -23.7645
14.4671 | -34.7471 | -23.773
14.4732 | -34.7476 | -23.7815
14.4793 | -34.748 | -23.79
14.4854 | -34.7485 | -23.7984
14.4915 | -34.749 | -23.8067
14.4976 | -34.7495 | -23.8151
14.5037 | -34.75 | -23.8235
14.5098 | -34.7505 | -23.8319
14.5159 | -34.751 | -23.8404
14.522 | -34.7514 | -23.8489
14.5281 | -34.7519 | -23.8573
14.5342 | -34.7524 | -23.8659
14.5403 | -34.7529 | -23.8742
14.5465 | -34.7534 | -23.8828
14.5526 | -34.7539 | -23.8914
14.5587 | -34.7544 | -23.8999
14.5648 | -34.7549 | -23.9083
14.5709 | -34.7554 | -23.9167
14.577 | -34.7559 | -23.925
14.5831 | -34.7563 | -23.9336
14.5892 | -34.7568 | -23.942
14.5953 | -34.7574 | -23.9504
14.6014 | -34.7578 | -23.9588
14.6075 | -34.7583 | -23.9672
14.6136 | -34.7587 | -23.9757
14.6197 | -34.7592 | -23.9841
14.6258 | -34.7597 | -23.9927
14.6319 | -34.7602 | -24.0012
14.638 | -34.7607 | -24.0097
14.6441 | -34.7612 | -24.0181
14.6502 | -34.7617 | -24.0267
14.6563 | -34.7622 | -24.0351
14.6624 | -34.7627 | -24.0436
14.6685 | -34.7632 | -24.0519
14.6746 | -34.7637 | -24.0604
14.6807 | -34.7642 | -24.0687
14.6869 | -34.7647 | -24.0772
14.693 | -34.7652 | -24.0856
14.6991 | -34.7656 | -24.094
14.7052 | -34.7662 | -24.1025
14.7113 | -34.7667 | -24.1111
14.7174 | -34.7672 | -24.1196
14.7235 | -34.7677 | -24.128
14.7296 | -34.7681 | -24.1363
14.7357 | -34.7686 | -24.1447
14.7418 | -34.7691 | -24.1532
14.7479 | -34.7696 | -24.1616
14.754 | -34.7701 | -24.1701
14.7601 | -34.7706 | -24.1785
14.7662 | -34.7711 | -24.1869
14.7723 | -34.7716 | -24.1953
14.7784 | -34.7721 | -24.2037
14.7845 | -34.7726 | -24.2121
14.7906 | -34.7731 | -24.2207
14.7967 | -34.7736 | -24.229
14.8028 | -34.7741 | -24.2374
14.8089 | -34.7746 | -24.2458
14.815 | -34.7751 | -24.2543
14.8211 | -34.7756 | -24.2628
14.8272 | -34.7761 | -24.2712
14.8334 | -34.7765 | -24.2796
14.8395 | -34.7771 | -24.288
14.8456 | -34.7776 | -24.2963
14.8517 | -34.7781 | -24.3048
14.8578 | -34.7786 | -24.3132
14.8639 | -34.7791 | -24.3218
14.87 | -34.7796 | -24.3302
14.8761 | -34.7801 | -24.3387
14.8822 | -34.7806 | -24.3471
14.8883 | -34.7811 | -24.3556
14.8944 | -34.7816 | -24.364
14.9005 | -34.7821 | -24.3725
14.9066 | -34.7826 | -24.3809
14.9127 | -34.7831 | -24.3893
14.9188 | -34.7836 | -24.3978
14.9249 | -34.7841 | -24.4063
14.931 | -34.7846 | -24.4148
14.9371 | -34.7851 | -24.4231
14.9432 | -34.7856 | -24.4317
14.9493 | -34.7861 | -24.4401
14.9554 | -34.7866 | -24.4485
14.9615 | -34.7871 | -24.457
14.9676 | -34.7876 | -24.4656
14.9738 | -34.7881 | -24.474
14.9799 | -34.7886 | -24.4826
14.986 | -34.7891 | -24.491
14.9921 | -34.7896 | -24.4994
14.9982 | -34.7901 | -24.5079
15.0043 | -34.7906 | -24.5162
15.0104 | -34.7911 | -24.5248
15.0165 | -34.7916 | -24.5333
15.0226 | -34.7921 | -24.5416
15.0287 | -34.7926 | -24.55
15.0348 | -34.7931 | -24.5584
15.0409 | -34.7936 | -24.5669
15.047 | -34.7941 | -24.5754
15.0531 | -34.7947 | -24.5837
15.0592 | -34.7952 | -24.5921
15.0653 | -34.7957 | -24.6007
15.0714 | -34.7961 | -24.609
15.0775 | -34.7966 | -24.6173
15.0836 | -34.7972 | -24.6258
15.0897 | -34.7977 | -24.6339
15.0958 | -34.7982 | -24.6424
15.1019 | -34.7987 | -24.6507
15.108 | -34.7992 | -24.659
15.1141 | -34.7997 | -24.6675
15.1203 | -34.8002 | -24.6758
15.1264 | -34.8007 | -24.6842
15.1325 | -34.8012 | -24.6927
15.1386 | -34.8017 | -24.7009
15.1447 | -34.8022 | -24.7093
15.1508 | -34.8028 | -24.7179
15.1569 | -34.8033 | -24.7262
15.163 | -34.8038 | -24.7344
15.1691 | -34.8043 | -24.7428
15.1752 | -34.8048 | -24.7513
15.1813 | -34.8053 | -24.7595
15.1874 | -34.8058 | -24.7679
15.1935 | -34.8063 | -24.776
15.1996 | -34.8068 | -24.7844
15.2057 | -34.8073 | -24.7928
15.2118 | -34.8078 | -24.8012
15.2179 | -34.8084 | -24.8097
15.224 | -34.8089 | -24.818
15.2301 | -34.8094 | -24.8261
15.2362 | -34.8099 | -24.8345
15.2423 | -34.8104 | -24.8429
15.2484 | -34.8109 | -24.8513
15.2545 | -34.8114 | -24.8596
15.2607 | -34.8119 | -24.8679
15.2668 | -34.8124 | -24.876
15.2729 | -34.8129 | -24.8843
15.279 | -34.8135 | -24.8928
15.2851 | -34.814 | -24.9012
15.2912 | -34.8145 | -24.9096
15.2973 | -34.815 | -24.9181
15.3034 | -34.8155 | -24.9262
15.3095 | -34.816 | -24.9346
15.3156 | -34.8166 | -24.9428
15.3217 | -34.8171 | -24.9511
15.3278 | -34.8176 | -24.9596
15.3339 | -34.8181 | -24.9678
15.34 | -34.8187 | -24.9761
15.3461 | -34.8192 | -24.9844
15.3522 | -34.8197 | -24.9926
15.3583 | -34.8202 | -25.001
15.3644 | -34.8207 | -25.0092
15.3705 | -34.8213 | -25.0172
15.3766 | -34.8217 | -25.0258
15.3827 | -34.8222 | -25.034
15.3888 | -34.8228 | -25.0422
15.3949 | -34.8233 | -25.0506
15.401 | -34.8238 | -25.0588
15.4072 | -34.8243 | -25.0669
15.4133 | -34.8248 | -25.0752
15.4194 | -34.8253 | -25.0837
15.4255 | -34.8258 | -25.0923
15.4316 | -34.8263 | -25.1003
15.4377 | -34.8268 | -25.1083
15.4438 | -34.8273 | -25.1165
15.4499 | -34.8278 | -25.1247
15.456 | -34.8283 | -25.1329
15.4621 | -34.8289 | -25.1411
15.4682 | -34.8294 | -25.1495
15.4743 | -34.8299 | -25.1576
15.4804 | -34.8304 | -25.1658
15.4865 | -34.8309 | -25.174
15.4926 | -34.8314 | -25.1824
15.4987 | -34.8319 | -25.1907
15.5048 | -34.8324 | -25.1991
15.5109 | -34.833 | -25.2071
15.517 | -34.8335 | -25.2155
15.5231 | -34.834 | -25.2238
15.5292 | -34.8345 | -25.232
15.5353 | -34.835 | -25.2404
15.5414 | -34.8355 | -25.2487
15.5476 | -34.836 | -25.2565
15.5537 | -34.8366 | -25.2647
15.5598 | -34.8371 | -25.2729
15.5659 | -34.8376 | -25.2812
15.572 | -34.8381 | -25.2895
15.5781 | -34.8386 | -25.2976
15.5842 | -34.8391 | -25.3057
15.5903 | -34.8397 | -25.3139
15.5964 | -34.8402 | -25.3223
15.6025 | -34.8407 | -25.3304
15.6086 | -34.8412 | -25.3387
15.6147 | -34.8417 | -25.3471
15.6208 | -34.8422 | -25.3551
15.6269 | -34.8428 | -25.3634
15.633 | -34.8433 | -25.3718
15.6391 | -34.8438 | -25.3802
15.6452 | -34.8443 | -25.3884
15.6513 | -34.8449 | -25.3962
15.6574 | -34.8454 | -25.4045
15.6635 | -34.8459 | -25.4127
15.6696 | -34.8465 | -25.4208
15.6757 | -34.847 | -25.429
15.6818 | -34.8475 | -25.4373
15.688 | -34.848 | -25.4452
15.6941 | -34.8485 | -25.4531
15.7002 | -34.8491 | -25.4612
15.7063 | -34.8496 | -25.4692
15.7124 | -34.8501 | -25.4776
15.7185 | -34.8506 | -25.4854
15.7246 | -34.8512 | -25.4935
15.7307 | -34.8517 | -25.5014
15.7368 | -34.8522 | -25.5093
15.7429 | -34.8528 | -25.5176
15.749 | -34.8533 | -25.5256
15.7551 | -34.8538 | -25.5338
15.7612 | -34.8544 | -25.5418
15.7673 | -34.8549 | -25.5497
15.7734 | -34.8554 | -25.5579
15.7795 | -34.856 | -25.5661
15.7856 | -34.8565 | -25.5742
15.7917 | -34.8571 | -25.5819
15.7978 | -34.8576 | -25.5903
15.8039 | -34.8581 | -25.5983
15.81 | -34.8586 | -25.6063
15.8161 | -34.8591 | -25.6141
15.8222 | -34.8596 | -25.6222
15.8283 | -34.8602 | -25.6299
15.8345 | -34.8607 | -25.638
15.8406 | -34.8612 | -25.6463
15.8467 | -34.8617 | -25.6545
15.8528 | -34.8622 | -25.6628
15.8589 | -34.8628 | -25.6706
15.865 | -34.8634 | -25.6783
15.8711 | -34.8639 | -25.6864
15.8772 | -34.8644 | -25.6944
15.8833 | -34.8649 | -25.7026
15.8894 | -34.8654 | -25.7104
15.8955 | -34.8659 | -25.7185
15.9016 | -34.8665 | -25.7261
15.9077 | -34.867 | -25.7341
15.9138 | -34.8675 | -25.7421
15.9199 | -34.868 | -25.7504
15.926 | -34.8686 | -25.7583
15.9321 | -34.8691 | -25.7661
15.9382 | -34.8696 | -25.7741
15.9443 | -34.8701 | -25.7821
15.9504 | -34.8706 | -25.79
15.9565 | -34.8712 | -25.798
15.9626 | -34.8717 | -25.8062
15.9687 | -34.8722 | -25.8139
15.9749 | -34.8727 | -25.8217
15.981 | -34.8732 | -25.8298
15.9871 | -34.8738 | -25.838
15.9932 | -34.8743 | -25.846
15.9993 | -34.8748 | -25.8543
16.0054 | -34.8754 | -25.8619
16.0115 | -34.8759 | -25.8701
16.0176 | -34.8764 | -25.8781
16.0237 | -34.8769 | -25.8865
16.0298 | -34.8774 | -25.8943
16.0359 | -34.8779 | -25.9023
16.042 | -34.8784 | -25.9101
16.0481 | -34.879 | -25.9181
16.0542 | -34.8795 | -25.9259
16.0603 | -34.88 | -25.9336
16.0664 | -34.8805 | -25.9411
16.0725 | -34.8811 | -25.9488
16.0786 | -34.8816 | -25.9563
16.0847 | -34.8821 | -25.9647
16.0908 | -34.8827 | -25.9728
16.0969 | -34.8832 | -25.9806
16.103 | -34.8838 | -25.9884
16.1091 | -34.8843 | -25.996
16.1152 | -34.8849 | -26.0041
16.1214 | -34.8854 | -26.012
16.1275 | -34.886 | -26.0202
16.1336 | -34.8865 | -26.0284
16.1397 | -34.8871 | -26.0365
16.1458 | -34.8876 | -26.0441
16.1519 | -34.8881 | -26.0523
16.158 | -34.8886 | -26.0605
16.1641 | -34.8891 | -26.0687
16.1702 | -34.8896 | -26.0766
16.1763 | -34.8902 | -26.0845
16.1824 | -34.8907 | -26.0921
16.1885 | -34.8912 | -26.1
16.1946 | -34.8917 | -26.1081
16.2007 | -34.8922 | -26.1159
16.2068 | -34.8928 | -26.124
16.2129 | -34.8933 | -26.132
16.219 | -34.8939 | -26.1396
16.2251 | -34.8944 | -26.1473
16.2312 | -34.8949 | -26.1551
16.2373 | -34.8954 | -26.1628
16.2434 | -34.896 | -26.1706
16.2495 | -34.8965 | -26.178
16.2556 | -34.897 | -26.1858
16.2618 | -34.8975 | -26.1936
16.2679 | -34.898 | -26.2014
16.274 | -34.8986 | -26.2091
16.2801 | -34.8991 | -26.217
16.2862 | -34.8997 | -26.2246
16.2923 | -34.9002 | -26.2327
16.2984 | -34.9007 | -26.2408
16.3045 | -34.9012 | -26.2486
16.3106 | -34.9017 | -26.2565
16.3167 | -34.9022 | -26.2642
16.3228 | -34.9028 | -26.2717
16.3289 | -34.9033 | -26.2795
16.335 | -34.9038 | -26.2874
16.3411 | -34.9044 | -26.2952
16.3472 | -34.9049 | -26.3032
16.3533 | -34.9054 | -26.3109
16.3594 | -34.906 | -26.3183
16.3655 | -34.9065 | -26.3263
16.3716 | -34.907 | -26.3341
16.3777 | -34.9075 | -26.3419
16.3838 | -34.908 | -26.3498
16.3899 | -34.9086 | -26.3573
16.396 | -34.9091 | -26.365
16.4021 | -34.9097 | -26.3726
16.4083 | -34.9102 | -26.3804
16.4144 | -34.9107 | -26.3883
16.4205 | -34.9112 | -26.3962
16.4266 | -34.9117 | -26.4037
16.4327 | -34.9122 | -26.4116
16.4388 | -34.9128 | -26.4194
16.4449 | -34.9133 | -26.427
16.451 | -34.9138 | -26.4348
16.4571 | -34.9143 | -26.4428
16.4632 | -34.9149 | -26.4504
16.4693 | -34.9154 | -26.4582
16.4754 | -34.9159 | -26.466
16.4815 | -34.9164 | -26.4737
16.4876 | -34.917 | -26.4814
16.4937 | -34.9174 | -26.4893
16.4998 | -34.918 | -26.4967
16.5059 | -34.9185 | -26.5045
16.512 | -34.919 | -26.5124
16.5181 | -34.9196 | -26.5201
16.5242 | -34.9201 | -26.5278
16.5303 | -34.9206 | -26.5353
16.5364 | -34.9211 | -26.5431
16.5425 | -34.9216 | -26.5508
16.5487 | -34.9222 | -26.5585
16.5548 | -34.9227 | -26.5663
16.5609 | -34.9233 | -26.5741
16.567 | -34.9238 | -26.5815
16.5731 | -34.9243 | -26.5894
16.5792 | -34.9249 | -26.5972
16.5853 | -34.9254 | -26.6051
16.5914 | -34.9259 | -26.6131
16.5975 | -34.9264 | -26.6209
16.6036 | -34.927 | -26.6284
16.6097 | -34.9275 | -26.6361
16.6158 | -34.9281 | -26.6438
16.6219 | -34.9286 | -26.6515
16.628 | -34.9291 | -26.6593
16.6341 | -34.9296 | -26.6672
16.6402 | -34.9301 | -26.6746
16.6463 | -34.9306 | -26.6823
16.6524 | -34.9311 | -26.6897
16.6585 | -34.9317 | -26.6976
16.6646 | -34.9322 | -26.7052
16.6707 | -34.9328 | -26.7128
16.6768 | -34.9333 | -26.7205
16.6829 | -34.9338 | -26.7283
16.689 | -34.9343 | -26.736
16.6952 | -34.9349 | -26.7436
16.7013 | -34.9354 | -26.7513
16.7074 | -34.9359 | -26.7588
16.7135 | -34.9365 | -26.7663
16.7196 | -34.9369 | -26.7741
16.7257 | -34.9375 | -26.7818
16.7318 | -34.938 | -26.7893
16.7379 | -34.9385 | -26.7969
16.744 | -34.939 | -26.8042
16.7501 | -34.9396 | -26.8121
16.7562 | -34.9401 | -26.8197
16.7623 | -34.9406 | -26.8273
16.7684 | -34.9411 | -26.835
16.7745 | -34.9416 | -26.8428
16.7806 | -34.9422 | -26.8502
16.7867 | -34.9427 | -26.8579
16.7928 | -34.9432 | -26.8654
16.7989 | -34.9437 | -26.8728
16.805 | -34.9443 | -26.8804
16.8111 | -34.9448 | -26.8877
16.8172 | -34.9453 | -26.8956
16.8233 | -34.9458 | -26.9032
16.8294 | -34.9463 | -26.9109
16.8356 | -34.9469 | -26.9186
16.8417 | -34.9474 | -26.9261
16.8478 | -34.9479 | -26.9335
16.8539 | -34.9484 | -26.9412
16.86 | -34.9489 | -26.9491
16.8661 | -34.9494 | -26.9567
16.8722 | -34.95 | -26.9643
16.8783 | -34.9505 | -26.972
16.8844 | -34.951 | -26.9792
16.8905 | -34.9516 | -26.9867
16.8966 | -34.9521 | -26.9943
16.9027 | -34.9526 | -27.002
16.9088 | -34.9531 | -27.0096
16.9149 | -34.9536 | -27.0172
16.921 | -34.9541 | -27.0245
16.9271 | -34.9547 | -27.0323
16.9332 | -34.9552 | -27.0399
16.9393 | -34.9557 | -27.0474
16.9454 | -34.9562 | -27.0549
16.9515 | -34.9568 | -27.0626
16.9576 | -34.9573 | -27.0698
16.9637 | -34.9578 | -27.0774
16.9698 | -34.9584 | -27.085
16.9759 | -34.9589 | -27.0925
16.9821 | -34.9594 | -27.1001
16.9882 | -34.96 | -27.1075
16.9943 | -34.9605 | -27.1149
17.0004 | -34.961 | -27.1224
17.0065 | -34.9616 | -27.1299
17.0126 | -34.9621 | -27.1373
17.0187 | -34.9626 | -27.1448
17.0248 | -34.9631 | -27.1523
17.0309 | -34.9636 | -27.1597
17.037 | -34.9642 | -27.1672
17.0431 | -34.9647 | -27.1747
17.0492 | -34.9652 | -27.1822
17.0553 | -34.9657 | -27.1898
17.0614 | -34.9662 | -27.1972
17.0675 | -34.9667 | -27.2046
17.0736 | -34.9672 | -27.2123
17.0797 | -34.9678 | -27.2199
17.0858 | -34.9683 | -27.2275
17.0919 | -34.9689 | -27.2349
17.098 | -34.9694 | -27.2423
17.1041 | -34.9699 | -27.2499
17.1102 | -34.9704 | -27.2575
17.1163 | -34.9709 | -27.2649
17.1225 | -34.9714 | -27.2723
17.1286 | -34.9719 | -27.2796
17.1347 | -34.9725 | -27.2871
17.1408 | -34.973 | -27.2945
17.1469 | -34.9735 | -27.302
17.153 | -34.974 | -27.3096
17.1591 | -34.9745 | -27.3173
17.1652 | -34.9751 | -27.3245
17.1713 | -34.9756 | -27.3318
17.1774 | -34.9761 | -27.3394
17.1835 | -34.9767 | -27.3469
17.1896 | -34.9772 | -27.3543
17.1957 | -34.9777 | -27.3619
17.2018 | -34.9782 | -27.3691
17.2079 | -34.9787 | -27.3765
17.214 | -34.9792 | -27.3841
17.2201 | -34.9798 | -27.3916
17.2262 | -34.9803 | -27.3993
17.2323 | -34.9808 | -27.4066
17.2384 | -34.9813 | -27.414
17.2445 | -34.9818 | -27.4214
17.2506 | -34.9824 | -27.4288
17.2567 | -34.9829 | -27.4363
17.2628 | -34.9834 | -27.4437
17.269 | -34.984 | -27.4509
17.2751 | -34.9845 | -27.4583
17.2812 | -34.985 | -27.4659
17.2873 | -34.9855 | -27.4735
17.2934 | -34.986 | -27.4809
17.2995 | -34.9866 | -27.4883
17.3056 | -34.9871 | -27.4954
17.3117 | -34.9876 | -27.503
17.3178 | -34.9881 | -27.5103
17.3239 | -34.9886 | -27.5177
17.33 | -34.9891 | -27.5252
17.3361 | -34.9896 | -27.5326
17.3422 | -34.9901 | -27.54
17.3483 | -34.9907 | -27.5476
17.3544 | -34.9911 | -27.5548
17.3605 | -34.9917 | -27.5623
17.3666 | -34.9922 | -27.5698
17.3727 | -34.9927 | -27.5773
17.3788 | -34.9932 | -27.5845
17.3849 | -34.9938 | -27.5919
17.391 | -34.9943 | -27.5994
17.3971 | -34.9948 | -27.6069
17.4032 | -34.9953 | -27.6144
17.4094 | -34.9959 | -27.6214
17.4155 | -34.9964 | -27.6288
17.4216 | -34.9969 | -27.6361
17.4277 | -34.9974 | -27.6437
17.4338 | -34.9979 | -27.651
17.4399 | -34.9984 | -27.6585
17.446 | -34.9989 | -27.6657
17.4521 | -34.9995 | -27.6732
17.4582 | -35 | -27.6804
17.4643 | -35.0005 | -27.6878
17.4704 | -35.001 | -27.6952
17.4765 | -35.0015 | -27.7028
17.4826 | -35.0021 | -27.7102
17.4887 | -35.0026 | -27.7177
17.4948 | -35.0031 | -27.7251
17.5009 | -35.0036 | -27.7324
17.507 | -35.0041 | -27.7397
17.5131 | -35.0046 | -27.7471
17.5192 | -35.0051 | -27.7542
17.5253 | -35.0056 | -27.7616
17.5314 | -35.0061 | -27.7691
17.5375 | -35.0066 | -27.7763
17.5436 | -35.0072 | -27.7838
17.5497 | -35.0077 | -27.7912
17.5559 | -35.0082 | -27.7983
17.562 | -35.0087 | -27.8058
17.5681 | -35.0092 | -27.8131
17.5742 | -35.0097 | -27.8206
17.5803 | -35.0102 | -27.8279
17.5864 | -35.0108 | -27.8351
17.5925 | -35.0113 | -27.8426
17.5986 | -35.0118 | -27.8497
17.6047 | -35.0123 | -27.8571
17.6108 | -35.0128 | -27.8645
17.6169 | -35.0133 | -27.8719
17.623 | -35.0138 | -27.8789
17.6291 | -35.0143 | -27.8863
17.6352 | -35.0148 | -27.8936
17.6413 | -35.0153 | -27.901
17.6474 | -35.0158 | -27.9085
17.6535 | -35.0163 | -27.9157
17.6596 | -35.0169 | -27.9231
17.6657 | -35.0174 | -27.9303
17.6718 | -35.0179 | -27.9376
17.6779 | -35.0184 | -27.9449
17.684 | -35.019 | -27.9524
17.6901 | -35.0195 | -27.9598
17.6963 | -35.02 | -27.9671
17.7024 | -35.0205 | -27.9744
17.7085 | -35.021 | -27.9816
17.7146 | -35.0215 | -27.9889
17.7207 | -35.022 | -27.9962
17.7268 | -35.0225 | -28.0034
17.7329 | -35.023 | -28.0108
17.739 | -35.0236 | -28.0183
17.7451 | -35.0241 | -28.0255
17.7512 | -35.0246 | -28.0327
17.7573 | -35.0251 | -28.0402
17.7634 | -35.0256 | -28.0472
17.7695 | -35.0261 | -28.0546
17.7756 | -35.0266 | -28.062
17.7817 | -35.0271 | -28.0693
17.7878 | -35.0277 | -28.0766
17.7939 | -35.0282 | -28.084
17.8 | -35.0287 | -28.0913
17.8061 | -35.0292 | -28.0985
17.8122 | -35.0297 | -28.1058
17.8183 | -35.0302 | -28.1132
17.8244 | -35.0308 | -28.1206
17.8305 | -35.0312 | -28.1278
17.8366 | -35.0317 | -28.1349
17.8428 | -35.0322 | -28.1423
17.8489 | -35.0328 | -28.1495
17.855 | -35.0333 | -28.1567
17.8611 | -35.0337 | -28.164
17.8672 | -35.0342 | -28.1711
17.8733 | -35.0347 | -28.1783
17.8794 | -35.0352 | -28.1854
17.8855 | -35.0358 | -28.1928
17.8916 | -35.0363 | -28.2
17.8977 | -35.0368 | -28.2074
17.9038 | -35.0373 | -28.2144
17.9099 | -35.0378 | -28.2218
17.916 | -35.0383 | -28.2292
17.9221 | -35.0388 | -28.2364
17.9282 | -35.0393 | -28.2436
17.9343 | -35.0398 | -28.251
17.9404 | -35.0403 | -28.2581
17.9465 | -35.0408 | -28.2655
17.9526 | -35.0413 | -28.2728
17.9587 | -35.0418 | -28.28
17.9648 | -35.0423 | -28.2872
17.9709 | -35.0428 | -28.2946
17.977 | -35.0433 | -28.3018
17.9832 | -35.0438 | -28.3091
17.9893 | -35.0443 | -28.3164
17.9954 | -35.0448 | -28.3238
18.0015 | -35.0453 | -28.331
18.0076 | -35.0458 | -28.3379
18.0137 | -35.0463 | -28.3452
18.0198 | -35.0468 | -28.3525
18.0259 | -35.0473 | -28.3597
18.032 | -35.0478 | -28.3671
18.0381 | -35.0484 | -28.3745
18.0442 | -35.0489 | -28.3819
18.0503 | -35.0494 | -28.3893
18.0564 | -35.0498 | -28.3965
18.0625 | -35.0503 | -28.4038
18.0686 | -35.0508 | -28.4111
18.0747 | -35.0513 | -28.4184
18.0808 | -35.0518 | -28.4255
18.0869 | -35.0523 | -28.4326
18.093 | -35.0529 | -28.4398
18.0991 | -35.0534 | -28.4473
18.1052 | -35.0539 | -28.4544
18.1113 | -35.0544 | -28.4618
18.1174 | -35.0549 | -28.4689
18.1236 | -35.0554 | -28.4763
18.1297 | -35.0559 | -28.4836
18.1358 | -35.0564 | -28.4909
18.1419 | -35.0569 | -28.4981
18.148 | -35.0574 | -28.5051
18.1541 | -35.0579 | -28.5122
18.1602 | -35.0584 | -28.5195
18.1663 | -35.0589 | -28.5269
18.1724 | -35.0594 | -28.5342
18.1785 | -35.0599 | -28.5417
18.1846 | -35.0604 | -28.5487
18.1907 | -35.0609 | -28.556
18.1968 | -35.0614 | -28.5631
18.2029 | -35.0619 | -28.5704
18.209 | -35.0624 | -28.5778
18.2151 | -35.0629 | -28.585
18.2212 | -35.0634 | -28.5921
18.2273 | -35.0639 | -28.5993
18.2334 | -35.0644 | -28.6066
18.2395 | -35.0649 | -28.6138
18.2456 | -35.0654 | -28.6211
18.2517 | -35.0659 | -28.6283
18.2578 | -35.0664 | -28.6355
18.2639 | -35.0669 | -28.6428
18.2701 | -35.0674 | -28.6499
18.2762 | -35.0679 | -28.6569
18.2823 | -35.0683 | -28.6641
18.2884 | -35.0688 | -28.6714
18.2945 | -35.0694 | -28.6785
18.3006 | -35.0699 | -28.6857
18.3067 | -35.0704 | -28.6929
18.3128 | -35.0709 | -28.7001
18.3189 | -35.0714 | -28.7073
18.325 | -35.0719 | -28.7146
18.3311 | -35.0724 | -28.7218
18.3372 | -35.0729 | -28.7289
18.3433 | -35.0734 | -28.7363
18.3494 | -35.0739 | -28.7435
18.3555 | -35.0744 | -28.7506
18.3616 | -35.0749 | -28.7576
18.3677 | -35.0755 | -28.7648
18.3738 | -35.076 | -28.7719
18.3799 | -35.0765 | -28.779
18.386 | -35.077 | -28.7863
18.3921 | -35.0774 | -28.7934
18.3982 | -35.0779 | -28.8005
18.4043 | -35.0784 | -28.8076
18.4105 | -35.0789 | -28.8147
18.4166 | -35.0794 | -28.8221
18.4227 | -35.0799 | -28.8293
18.4288 | -35.0804 | -28.8366
18.4349 | -35.0809 | -28.8437
18.441 | -35.0814 | -28.8509
18.4471 | -35.0818 | -28.8581
18.4532 | -35.0823 | -28.8654
18.4593 | -35.0828 | -28.8727
18.4654 | -35.0833 | -28.8797
18.4715 | -35.0838 | -28.8871
18.4776 | -35.0843 | -28.8945
18.4837 | -35.0848 | -28.9019
18.4898 | -35.0853 | -28.9092
18.4959 | -35.0858 | -28.9164
18.502 | -35.0863 | -28.9236
18.5081 | -35.0868 | -28.9307
18.5142 | -35.0873 | -28.9379
18.5203 | -35.0878 | -28.9451
18.5264 | -35.0883 | -28.9521
18.5325 | -35.0888 | -28.9594
18.5386 | -35.0894 | -28.9665
18.5447 | -35.0898 | -28.9739
18.5508 | -35.0903 | -28.9811
18.557 | -35.0908 | -28.9884
18.5631 | -35.0913 | -28.9956
18.5692 | -35.0918 | -29.0028
18.5753 | -35.0923 | -29.0098
18.5814 | -35.0928 | -29.017
18.5875 | -35.0933 | -29.0241
18.5936 | -35.0938 | -29.0312
18.5997 | -35.0942 | -29.0383
18.6058 | -35.0947 | -29.0452
18.6119 | -35.0953 | -29.0522
18.618 | -35.0958 | -29.0593
18.6241 | -35.0963 | -29.0665
18.6302 | -35.0967 | -29.0739
18.6363 | -35.0973 | -29.081
18.6424 | -35.0978 | -29.0882
18.6485 | -35.0982 | -29.0955
18.6546 | -35.0987 | -29.1025
18.6607 | -35.0992 | -29.1096
18.6668 | -35.0997 | -29.1168
18.6729 | -35.1002 | -29.1239
18.679 | -35.1008 | -29.131
18.6851 | -35.1013 | -29.1379
18.6912 | -35.1017 | -29.1452
18.6974 | -35.1022 | -29.1523
18.7035 | -35.1027 | -29.1596
18.7096 | -35.1032 | -29.167
18.7157 | -35.1038 | -29.1741
18.7218 | -35.1042 | -29.1812
18.7279 | -35.1047 | -29.1885
18.734 | -35.1052 | -29.1956
18.7401 | -35.1057 | -29.2028
18.7462 | -35.1062 | -29.2096
18.7523 | -35.1067 | -29.2167
18.7584 | -35.1072 | -29.2241
18.7645 | -35.1077 | -29.2314
18.7706 | -35.1082 | -29.2385
18.7767 | -35.1087 | -29.2458
18.7828 | -35.1092 | -29.2531
18.7889 | -35.1097 | -29.2602
18.795 | -35.1102 | -29.2675
18.8011 | -35.1107 | -29.2746
18.8072 | -35.1112 | -29.282
18.8133 | -35.1116 | -29.289
18.8194 | -35.1121 | -29.2959
18.8255 | -35.1126 | -29.3031
18.8316 | -35.1131 | -29.3102
18.8377 | -35.1136 | -29.3173
18.8439 | -35.1141 | -29.3245
18.85 | -35.1146 | -29.3316
18.8561 | -35.1151 | -29.3388
18.8622 | -35.1156 | -29.346
18.8683 | -35.1161 | -29.3531
18.8744 | -35.1166 | -29.36
18.8805 | -35.1171 | -29.3672
18.8866 | -35.1175 | -29.3744
18.8927 | -35.1181 | -29.3815
18.8988 | -35.1186 | -29.3886
18.9049 | -35.1191 | -29.3959
18.911 | -35.1195 | -29.403
18.9171 | -35.12 | -29.4098
18.9232 | -35.1205 | -29.4169
18.9293 | -35.121 | -29.4242
18.9354 | -35.1216 | -29.4312
18.9415 | -35.1221 | -29.4384
18.9476 | -35.1225 | -29.4455
18.9537 | -35.123 | -29.4526
18.9598 | -35.1235 | -29.4596
18.9659 | -35.124 | -29.4666
18.972 | -35.1245 | -29.4735
18.9781 | -35.125 | -29.4807
18.9843 | -35.1254 | -29.4878
18.9904 | -35.1259 | -29.4952
18.9965 | -35.1264 | -29.5022
19.0026 | -35.1269 | -29.5092
19.0087 | -35.1274 | -29.5165
19.0148 | -35.1279 | -29.5237
19.0209 | -35.1284 | -29.5309
19.027 | -35.1289 | -29.5382
19.0331 | -35.1294 | -29.5453
19.0392 | -35.1299 | -29.5526
19.0453 | -35.1304 | -29.5595
19.0514 | -35.1309 | -29.5668
19.0575 | -35.1314 | -29.5741
19.0636 | -35.1319 | -29.5808
19.0697 | -35.1324 | -29.5876
19.0758 | -35.1328 | -29.5946
19.0819 | -35.1333 | -29.6016
19.088 | -35.1338 | -29.6087
19.0941 | -35.1343 | -29.6157
19.1002 | -35.1348 | -29.6227
19.1063 | -35.1353 | -29.63
19.1124 | -35.1358 | -29.6372
19.1185 | -35.1363 | -29.6443
19.1246 | -35.1368 | -29.6515
19.1308 | -35.1373 | -29.6586
19.1369 | -35.1378 | -29.6656
19.143 | -35.1383 | -29.6726
19.1491 | -35.1388 | -29.6797
19.1552 | -35.1393 | -29.6869
19.1613 | -35.1398 | -29.6942
19.1674 | -35.1403 | -29.7013
19.1735 | -35.1407 | -29.7083
19.1796 | -35.1413 | -29.7154
19.1857 | -35.1418 | -29.7226
19.1918 | -35.1422 | -29.7298
19.1979 | -35.1427 | -29.737
19.204 | -35.1433 | -29.7441
19.2101 | -35.1438 | -29.7511
19.2162 | -35.1442 | -29.7581
19.2223 | -35.1447 | -29.7654
19.2284 | -35.1452 | -29.7723
19.2345 | -35.1457 | -29.7793
19.2406 | -35.1462 | -29.7863
19.2467 | -35.1467 | -29.7932
19.2528 | -35.1472 | -29.8004
19.2589 | -35.1477 | -29.8076
19.265 | -35.1482 | -29.8149
19.2712 | -35.1487 | -29.822
19.2773 | -35.1492 | -29.8289
19.2834 | -35.1497 | -29.8359
19.2895 | -35.1502 | -29.8429
19.2956 | -35.1507 | -29.8501
19.3017 | -35.1511 | -29.8572
19.3078 | -35.1516 | -29.8642
19.3139 | -35.1521 | -29.8713
19.32 | -35.1526 | -29.8783
19.3261 | -35.1531 | -29.8855
19.3322 | -35.1535 | -29.8924
19.3383 | -35.154 | -29.8994
19.3444 | -35.1545 | -29.9065
19.3505 | -35.155 | -29.9138
19.3566 | -35.1555 | -29.9208
19.3627 | -35.156 | -29.9276
19.3688 | -35.1565 | -29.9346
19.3749 | -35.157 | -29.9416
19.381 | -35.1575 | -29.9485
19.3871 | -35.158 | -29.9557
19.3932 | -35.1585 | -29.9625
19.3993 | -35.1589 | -29.9698
19.4054 | -35.1594 | -29.977
19.4115 | -35.1599 | -29.9841
19.4177 | -35.1604 | -29.991
19.4238 | -35.1609 | -29.9981
19.4299 | -35.1614 | -30.0051
19.436 | -35.1619 | -30.0124
19.4421 | -35.1624 | -30.0195
19.4482 | -35.1629 | -30.0264
19.4543 | -35.1634 | -30.0333
19.4604 | -35.1639 | -30.0402
19.4665 | -35.1644 | -30.0472
19.4726 | -35.1648 | -30.0542
19.4787 | -35.1653 | -30.0613
19.4848 | -35.1658 | -30.0684
19.4909 | -35.1663 | -30.0754
19.497 | -35.1668 | -30.0826
19.5031 | -35.1673 | -30.0895
19.5092 | -35.1678 | -30.0965
19.5153 | -35.1682 | -30.1034
19.5214 | -35.1687 | -30.1105
19.5275 | -35.1692 | -30.1175
19.5336 | -35.1697 | -30.1245
19.5397 | -35.1702 | -30.1314
19.5458 | -35.1706 | -30.1385
19.5519 | -35.1711 | -30.1456
19.5581 | -35.1716 | -30.1524
19.5642 | -35.1721 | -30.1594
19.5703 | -35.1725 | -30.1666
19.5764 | -35.173 | -30.1737
19.5825 | -35.1735 | -30.1806
19.5886 | -35.174 | -30.1878
19.5947 | -35.1745 | -30.1947
19.6008 | -35.175 | -30.2017
19.6069 | -35.1755 | -30.2086
19.613 | -35.176 | -30.2156
19.6191 | -35.1765 | -30.2225
19.6252 | -35.177 | -30.2296
19.6313 | -35.1774 | -30.2362
19.6374 | -35.1779 | -30.2433
19.6435 | -35.1784 | -30.2502
19.6496 | -35.1789 | -30.2572
19.6557 | -35.1794 | -30.2643
19.6618 | -35.1799 | -30.2713
19.6679 | -35.1804 | -30.2783
19.674 | -35.1809 | -30.2854
19.6801 | -35.1814 | -30.2924
19.6862 | -35.1818 | -30.2994
19.6923 | -35.1823 | -30.3065
19.6984 | -35.1829 | -30.3135
19.7046 | -35.1834 | -30.3204
19.7107 | -35.1839 | -30.3274
19.7168 | -35.1844 | -30.3343
19.7229 | -35.1849 | -30.3415
19.729 | -35.1854 | -30.3483
19.7351 | -35.1859 | -30.3549
19.7412 | -35.1864 | -30.3619
19.7473 | -35.1868 | -30.3689
19.7534 | -35.1873 | -30.3761
19.7595 | -35.1878 | -30.3829
19.7656 | -35.1883 | -30.39
19.7717 | -35.1888 | -30.3966
19.7778 | -35.1892 | -30.4035
19.7839 | -35.1897 | -30.4104
19.79 | -35.1902 | -30.4174
19.7961 | -35.1907 | -30.4245
19.8022 | -35.1911 | -30.4312
19.8083 | -35.1916 | -30.4383
19.8144 | -35.1921 | -30.4453
19.8205 | -35.1926 | -30.4523
19.8266 | -35.1931 | -30.4592
19.8327 | -35.1936 | -30.466
19.8388 | -35.1941 | -30.4728
19.845 | -35.1945 | -30.4798
19.8511 | -35.195 | -30.4867
19.8572 | -35.1955 | -30.4934
19.8633 | -35.1959 | -30.5003
19.8694 | -35.1964 | -30.5074
19.8755 | -35.1969 | -30.5142
19.8816 | -35.1974 | -30.5213
19.8877 | -35.1979 | -30.5282
19.8938 | -35.1984 | -30.5353
19.8999 | -35.1989 | -30.5423
19.906 | -35.1994 | -30.5492
19.9121 | -35.1999 | -30.5557
19.9182 | -35.2003 | -30.5629
19.9243 | -35.2008 | -30.5699
19.9304 | -35.2013 | -30.5768
19.9365 | -35.2018 | -30.5838
19.9426 | -35.2023 | -30.5907
19.9487 | -35.2027 | -30.5976
19.9548 | -35.2032 | -30.6046
19.9609 | -35.2037 | -30.6117
19.967 | -35.2042 | -30.6186
19.9731 | -35.2047 | -30.6255
19.9792 | -35.2051 | -30.6322
19.9853 | -35.2056 | -30.6389
19.9915 | -35.2061 | -30.6458
19.9976 | -35.2066 | -30.6527
20.0037 | -35.2071 | -30.6594
20.0098 | -35.2076 | -30.6663
20.0159 | -35.2081 | -30.673
20.022 | -35.2085 | -30.6799
20.0281 | -35.209 | -30.687
20.0342 | -35.2095 | -30.6941
20.0403 | -35.21 | -30.7009
20.0464 | -35.2105 | -30.7077
20.0525 | -35.211 | -30.7146
20.0586 | -35.2115 | -30.7217
20.0647 | -35.2119 | -30.7287
20.0708 | -35.2124 | -30.7357
20.0769 | -35.2129 | -30.7427
20.083 | -35.2134 | -30.7495
20.0891 | -35.2139 | -30.7564
20.0952 | -35.2144 | -30.7633
20.1013 | -35.2148 | -30.7704
20.1074 | -35.2154 | -30.7776
20.1135 | -35.2158 | -30.7848
20.1196 | -35.2163 | -30.7917
20.1257 | -35.2168 | -30.7987
20.1319 | -35.2173 | -30.8056
20.138 | -35.2177 | -30.8127
20.1441 | -35.2182 | -30.8194
20.1502 | -35.2187 | -30.8262
20.1563 | -35.2192 | -30.8328
20.1624 | -35.2197 | -30.8396
20.1685 | -35.2201 | -30.8465
20.1746 | -35.2206 | -30.8533
20.1807 | -35.221 | -30.8603
20.1868 | -35.2215 | -30.8672
20.1929 | -35.222 | -30.874
20.199 | -35.2225 | -30.8808
20.2051 | -35.223 | -30.8876
20.2112 | -35.2234 | -30.8945
20.2173 | -35.2239 | -30.9013
20.2234 | -35.2244 | -30.9082
20.2295 | -35.2249 | -30.915
20.2356 | -35.2254 | -30.9221
20.2417 | -35.2259 | -30.9289
20.2478 | -35.2263 | -30.9359
20.2539 | -35.2268 | -30.9429
20.26 | -35.2273 | -30.9497
20.2661 | -35.2277 | -30.9566
20.2723 | -35.2282 | -30.9636
20.2784 | -35.2287 | -30.9706
20.2845 | -35.2292 | -30.9774
20.2906 | -35.2297 | -30.9846
20.2967 | -35.2302 | -30.9911
20.3028 | -35.2307 | -30.9977
20.3089 | -35.2311 | -31.0045
20.315 | -35.2317 | -31.011
20.3211 | -35.2321 | -31.0177
20.3272 | -35.2326 | -31.0246
20.3333 | -35.2332 | -31.0309
20.3394 | -35.2336 | -31.0376
20.3455 | -35.2341 | -31.0447
20.3516 | -35.2346 | -31.0517
20.3577 | -35.2351 | -31.0586
20.3638 | -35.2356 | -31.0652
20.3699 | -35.236 | -31.0717
20.376 | -35.2365 | -31.0788
20.3821 | -35.237 | -31.0859
20.3882 | -35.2375 | -31.0928
20.3943 | -35.238 | -31.0996
20.4004 | -35.2385 | -31.1062
20.4065 | -35.2389 | -31.113
20.4126 | -35.2394 | -31.1198
20.4188 | -35.2399 | -31.1267
20.4249 | -35.2404 | -31.1336
20.431 | -35.2409 | -31.1404
20.4371 | -35.2414 | -31.147
20.4432 | -35.2418 | -31.1538
20.4493 | -35.2424 | -31.1606
20.4554 | -35.2428 | -31.1675
20.4615 | -35.2433 | -31.1742
20.4676 | -35.2438 | -31.181
20.4737 | -35.2443 | -31.1877
20.4798 | -35.2449 | -31.1946
20.4859 | -35.2453 | -31.2014
20.492 | -35.2458 | -31.2082
20.4981 | -35.2463 | -31.2148
20.5042 | -35.2468 | -31.2217
20.5103 | -35.2472 | -31.2278
20.5164 | -35.2477 | -31.2349
20.5225 | -35.2482 | -31.2417
20.5286 | -35.2487 | -31.2486
20.5347 | -35.2491 | -31.2551
20.5408 | -35.2496 | -31.2613
20.5469 | -35.2501 | -31.2681
20.553 | -35.2505 | -31.2748
20.5592 | -35.251 | -31.2818
20.5653 | -35.2515 | -31.2889
20.5714 | -35.252 | -31.2956
20.5775 | -35.2524 | -31.3022
20.5836 | -35.253 | -31.3089
20.5897 | -35.2534 | -31.3155
20.5958 | -35.2539 | -31.3221
20.6019 | -35.2544 | -31.3293
20.608 | -35.2549 | -31.3362
20.6141 | -35.2554 | -31.3427
20.6202 | -35.2559 | -31.3498
20.6263 | -35.2564 | -31.3569
20.6324 | -35.2568 | -31.3638
20.6385 | -35.2573 | -31.3704
20.6446 | -35.2578 | -31.3769
20.6507 | -35.2582 | -31.3833
20.6568 | -35.2587 | -31.3899
20.6629 | -35.2592 | -31.3971
20.669 | -35.2597 | -31.4036
20.6751 | -35.2602 | -31.4102
20.6812 | -35.2607 | -31.4166
20.6873 | -35.2611 | -31.4236
20.6934 | -35.2616 | -31.4303
20.6995 | -35.2621 | -31.4367
20.7057 | -35.2626 | -31.4435
20.7118 | -35.263 | -31.4505
20.7179 | -35.2635 | -31.4569
20.724 | -35.264 | -31.4639
20.7301 | -35.2645 | -31.4707
20.7362 | -35.265 | -31.4774
20.7423 | -35.2656 | -31.484
20.7484 | -35.2661 | -31.4908
20.7545 | -35.2666 | -31.4974
20.7606 | -35.2671 | -31.5042
20.7667 | -35.2676 | -31.511
20.7728 | -35.268 | -31.5175
20.7789 | -35.2685 | -31.5245
20.785 | -35.269 | -31.5312
20.7911 | -35.2695 | -31.5375
20.7972 | -35.2699 | -31.5443
20.8033 | -35.2705 | -31.5511
20.8094 | -35.2709 | -31.5576
20.8155 | -35.2714 | -31.5641
20.8216 | -35.2719 | -31.5708
20.8277 | -35.2724 | -31.5777
20.8338 | -35.2728 | -31.5845
20.8399 | -35.2733 | -31.5913
20.8461 | -35.2738 | -31.5982
20.8522 | -35.2743 | -31.605
20.8583 | -35.2748 | -31.6112
20.8644 | -35.2753 | -31.6178
20.8705 | -35.2758 | -31.6244
20.8766 | -35.2762 | -31.6312
20.8827 | -35.2767 | -31.6379
20.8888 | -35.2772 | -31.6446
20.8949 | -35.2777 | -31.651
20.901 | -35.2782 | -31.6576
20.9071 | -35.2787 | -31.664
20.9132 | -35.2792 | -31.6708
20.9193 | -35.2796 | -31.6777
20.9254 | -35.2801 | -31.6845
20.9315 | -35.2807 | -31.6908
20.9376 | -35.2811 | -31.6975
20.9437 | -35.2816 | -31.7041
20.9498 | -35.2821 | -31.7109
20.9559 | -35.2826 | -31.7177
20.962 | -35.283 | -31.7246
20.9681 | -35.2836 | -31.7311
20.9742 | -35.2841 | -31.7379
20.9803 | -35.2845 | -31.7447
20.9864 | -35.285 | -31.7514
20.9926 | -35.2854 | -31.758
20.9987 | -35.286 | -31.7648
21.0048 | -35.2864 | -31.7714
21.0109 | -35.2869 | -31.7781
21.017 | -35.2874 | -31.7848
21.0231 | -35.2878 | -31.7912
21.0292 | -35.2884 | -31.798
21.0353 | -35.2888 | -31.8047
21.0414 | -35.2892 | -31.8113
21.0475 | -35.2897 | -31.818
21.0536 | -35.2902 | -31.8247
21.0597 | -35.2907 | -31.8315
21.0658 | -35.2912 | -31.8382
21.0719 | -35.2917 | -31.8443
21.078 | -35.2922 | -31.8511
21.0841 | -35.2927 | -31.8578
21.0902 | -35.2931 | -31.8643
21.0963 | -35.2935 | -31.8705
21.1024 | -35.294 | -31.8772
21.1085 | -35.2944 | -31.8833
21.1146 | -35.2949 | -31.8896
21.1207 | -35.2954 | -31.8963
21.1268 | -35.2959 | -31.9032
21.133 | -35.2964 | -31.9098
21.1391 | -35.2969 | -31.9161
21.1452 | -35.2974 | -31.9224
21.1513 | -35.2979 | -31.9293
21.1574 | -35.2984 | -31.9362
21.1635 | -35.2988 | -31.9426
21.1696 | -35.2993 | -31.9488
21.1757 | -35.2998 | -31.9552
21.1818 | -35.3003 | -31.9621
21.1879 | -35.3007 | -31.9688
21.194 | -35.3012 | -31.9755
21.2001 | -35.3017 | -31.9818
21.2062 | -35.3022 | -31.9885
21.2123 | -35.3026 | -31.9949
21.2184 | -35.3031 | -32.0012
21.2245 | -35.3036 | -32.0079
21.2306 | -35.3041 | -32.0143
21.2367 | -35.3045 | -32.021
21.2428 | -35.3051 | -32.0273
21.2489 | -35.3055 | -32.0338
21.255 | -35.3059 | -32.0404
21.2611 | -35.3064 | -32.047
21.2672 | -35.3069 | -32.0536
21.2733 | -35.3074 | -32.0601
21.2795 | -35.3078 | -32.0665
21.2856 | -35.3083 | -32.073
21.2917 | -35.3088 | -32.0795
21.2978 | -35.3092 | -32.0859
21.3039 | -35.3097 | -32.0924
21.31 | -35.3102 | -32.099
21.3161 | -35.3106 | -32.1053
21.3222 | -35.3111 | -32.1121
21.3283 | -35.3116 | -32.1185
21.3344 | -35.3121 | -32.1248
21.3405 | -35.3126 | -32.1315
21.3466 | -35.3131 | -32.1379
21.3527 | -35.3136 | -32.1442
21.3588 | -35.3141 | -32.1508
21.3649 | -35.3146 | -32.1575
21.371 | -35.3151 | -32.164
21.3771 | -35.3156 | -32.1704
21.3832 | -35.3161 | -32.177
21.3893 | -35.3166 | -32.1833
21.3954 | -35.317 | -32.1898
21.4015 | -35.3175 | -32.1963
21.4076 | -35.318 | -32.2031
21.4137 | -35.3185 | -32.2096
21.4199 | -35.3189 | -32.2159
21.426 | -35.3194 | -32.2219
21.4321 | -35.3199 | -32.2284
21.4382 | -35.3203 | -32.2346
21.4443 | -35.3208 | -32.2412
21.4504 | -35.3213 | -32.248
21.4565 | -35.3218 | -32.2544
21.4626 | -35.3223 | -32.2611
21.4687 | -35.3228 | -32.2675
21.4748 | -35.3232 | -32.274
21.4809 | -35.3237 | -32.2806
21.487 | -35.3242 | -32.2871
21.4931 | -35.3247 | -32.2935
21.4992 | -35.3251 | -32.3
21.5053 | -35.3256 | -32.3065
21.5114 | -35.3261 | -32.3128
21.5175 | -35.3265 | -32.3194
21.5236 | -35.327 | -32.3258
21.5297 | -35.3275 | -32.3323
21.5358 | -35.328 | -32.3389
21.5419 | -35.3285 | -32.3456
21.548 | -35.3289 | -32.3523
21.5541 | -35.3294 | -32.3587
21.5602 | -35.3298 | -32.3652
21.5664 | -35.3303 | -32.3715
21.5725 | -35.3308 | -32.3779
21.5786 | -35.3313 | -32.3843
21.5847 | -35.3317 | -32.3909
21.5908 | -35.3322 | -32.3973
21.5969 | -35.3327 | -32.4035
21.603 | -35.3331 | -32.4104
21.6091 | -35.3336 | -32.4171
21.6152 | -35.3341 | -32.4236
21.6213 | -35.3345 | -32.4301
21.6274 | -35.335 | -32.4367
21.6335 | -35.3355 | -32.4429
21.6396 | -35.3359 | -32.4495
21.6457 | -35.3364 | -32.4561
21.6518 | -35.3369 | -32.4627
21.6579 | -35.3374 | -32.4691
21.664 | -35.3379 | -32.4758
21.6701 | -35.3384 | -32.4818
21.6762 | -35.3389 | -32.4883
21.6823 | -35.3394 | -32.4951
21.6884 | -35.3399 | -32.5016
21.6945 | -35.3404 | -32.508
21.7006 | -35.3409 | -32.5146
21.7068 | -35.3413 | -32.521
21.7129 | -35.3418 | -32.5274
21.719 | -35.3423 | -32.5337
21.7251 | -35.3428 | -32.54
21.7312 | -35.3432 | -32.5464
21.7373 | -35.3437 | -32.5527
21.7434 | -35.3442 | -32.5593
21.7495 | -35.3446 | -32.5658
21.7556 | -35.3451 | -32.5722
21.7617 | -35.3456 | -32.5788
21.7678 | -35.346 | -32.5851
21.7739 | -35.3465 | -32.5911
21.78 | -35.347 | -32.5976
21.7861 | -35.3475 | -32.6041
21.7922 | -35.348 | -32.6105
21.7983 | -35.3484 | -32.6168
21.8044 | -35.3489 | -32.6233
21.8105 | -35.3494 | -32.6294
21.8166 | -35.3499 | -32.6358
21.8227 | -35.3503 | -32.6422
21.8288 | -35.3508 | -32.6488
21.8349 | -35.3512 | -32.6554
21.841 | -35.3517 | -32.6619
21.8471 | -35.3522 | -32.668
21.8533 | -35.3527 | -32.6745
21.8594 | -35.3531 | -32.681
21.8655 | -35.3536 | -32.6875
21.8716 | -35.354 | -32.6943
21.8777 | -35.3545 | -32.7006
21.8838 | -35.355 | -32.7071
21.8899 | -35.3555 | -32.7136
21.896 | -35.3559 | -32.7197
21.9021 | -35.3564 | -32.7261
21.9082 | -35.3569 | -32.7325
21.9143 | -35.3574 | -32.7388
21.9204 | -35.3578 | -32.7449
21.9265 | -35.3583 | -32.7513
21.9326 | -35.3588 | -32.758
21.9387 | -35.3593 | -32.7644
21.9448 | -35.3598 | -32.7708
21.9509 | -35.3602 | -32.7774
21.957 | -35.3606 | -32.784
21.9631 | -35.3611 | -32.7908
21.9692 | -35.3616 | -32.7974
21.9753 | -35.362 | -32.8037
21.9814 | -35.3625 | -32.8104
21.9875 | -35.3629 | -32.8163
21.9937 | -35.3634 | -32.8228
21.9998 | -35.3639 | -32.8291
22.0059 | -35.3645 | -32.8356
22.012 | -35.365 | -32.8421
22.0181 | -35.3654 | -32.8485
22.0242 | -35.3659 | -32.8549
22.0303 | -35.3664 | -32.8615
22.0364 | -35.3669 | -32.8678
22.0425 | -35.3673 | -32.874
22.0486 | -35.3677 | -32.8802
22.0547 | -35.3681 | -32.8861
22.0608 | -35.3686 | -32.8918
22.0669 | -35.369 | -32.898
22.073 | -35.3695 | -32.9039
22.0791 | -35.37 | -32.9099
22.0852 | -35.3705 | -32.9162
22.0913 | -35.3709 | -32.9221
22.0974 | -35.3714 | -32.9283
22.1035 | -35.3718 | -32.9345
22.1096 | -35.3723 | -32.9411
22.1157 | -35.3728 | -32.9476
22.1218 | -35.3732 | -32.9538
22.1279 | -35.3738 | -32.9601
22.134 | -35.3743 | -32.9665
22.1402 | -35.3747 | -32.9727
22.1463 | -35.3752 | -32.9793
22.1524 | -35.3757 | -32.9856
22.1585 | -35.3761 | -32.9917
22.1646 | -35.3766 | -32.9979
22.1707 | -35.3771 | -33.0043
22.1768 | -35.3776 | -33.0106
22.1829 | -35.3781 | -33.017
22.189 | -35.3786 | -33.0234
22.1951 | -35.379 | -33.0294
22.2012 | -35.3795 | -33.0356
22.2073 | -35.38 | -33.0417
22.2134 | -35.3805 | -33.0481
22.2195 | -35.3809 | -33.0543
22.2256 | -35.3814 | -33.0605
22.2317 | -35.3819 | -33.0668
22.2378 | -35.3823 | -33.073
22.2439 | -35.3829 | -33.0797
22.25 | -35.3833 | -33.0862
22.2561 | -35.3838 | -33.0925
22.2622 | -35.3842 | -33.0987
22.2683 | -35.3847 | -33.1048
22.2744 | -35.3851 | -33.1113
22.2806 | -35.3857 | -33.1175
22.2867 | -35.3861 | -33.1235
22.2928 | -35.3865 | -33.13
22.2989 | -35.387 | -33.1361
22.305 | -35.3875 | -33.1422
22.3111 | -35.3879 | -33.149
22.3172 | -35.3884 | -33.1556
22.3233 | -35.3889 | -33.1617
22.3294 | -35.3894 | -33.168
22.3355 | -35.3899 | -33.1738
22.3416 | -35.3904 | -33.1801
22.3477 | -35.3909 | -33.1864
22.3538 | -35.3913 | -33.1928
22.3599 | -35.3917 | -33.1993
22.366 | -35.3922 | -33.2053
22.3721 | -35.3927 | -33.2115
22.3782 | -35.3931 | -33.2181
22.3843 | -35.3936 | -33.2246
22.3904 | -35.394 | -33.2306
22.3965 | -35.3945 | -33.2372
22.4026 | -35.395 | -33.2433
22.4087 | -35.3955 | -33.2493
22.4148 | -35.396 | -33.2555
22.4209 | -35.3964 | -33.2617
22.4271 | -35.3969 | -33.2678
22.4332 | -35.3973 | -33.2741
22.4393 | -35.3978 | -33.2803
22.4454 | -35.3983 | -33.2865
22.4515 | -35.3987 | -33.2928
22.4576 | -35.3992 | -33.2991
22.4637 | -35.3997 | -33.3055
22.4698 | -35.4002 | -33.3117
22.4759 | -35.4007 | -33.3177
22.482 | -35.4011 | -33.3242
22.4881 | -35.4016 | -33.3304
22.4942 | -35.4021 | -33.3369
22.5003 | -35.4025 | -33.3433
22.5064 | -35.403 | -33.3495
22.5125 | -35.4035 | -33.3555
22.5186 | -35.404 | -33.3618
22.5247 | -35.4044 | -33.3682
22.5308 | -35.4049 | -33.3746
22.5369 | -35.4053 | -33.3808
22.543 | -35.4058 | -33.3871
22.5491 | -35.4062 | -33.3933
22.5552 | -35.4067 | -33.3995
22.5613 | -35.4072 | -33.406
22.5675 | -35.4076 | -33.4121
22.5736 | -35.4081 | -33.4184
22.5797 | -35.4086 | -33.4247
22.5858 | -35.409 | -33.4308
22.5919 | -35.4095 | -33.437
22.598 | -35.41 | -33.4434
22.6041 | -35.4104 | -33.4496
22.6102 | -35.4109 | -33.456
22.6163 | -35.4113 | -33.4622
22.6224 | -35.4118 | -33.4684
22.6285 | -35.4123 | -33.4747
22.6346 | -35.4127 | -33.4808
22.6407 | -35.4132 | -33.4872
22.6468 | -35.4137 | -33.4935
22.6529 | -35.4141 | -33.4997
22.659 | -35.4145 | -33.5059
22.6651 | -35.415 | -33.5121
22.6712 | -35.4154 | -33.5183
22.6773 | -35.4159 | -33.5245
22.6834 | -35.4164 | -33.5308
22.6895 | -35.4168 | -33.537
22.6956 | -35.4173 | -33.5432
22.7017 | -35.4177 | -33.5494
22.7079 | -35.4182 | -33.5557
22.714 | -35.4186 | -33.5617
22.7201 | -35.4191 | -33.568
22.7262 | -35.4195 | -33.5742
22.7323 | -35.42 | -33.5803
22.7384 | -35.4204 | -33.5866
22.7445 | -35.4209 | -33.593
22.7506 | -35.4214 | -33.5991
22.7567 | -35.4219 | -33.6053
22.7628 | -35.4223 | -33.6112
22.7689 | -35.4227 | -33.6174
22.775 | -35.4232 | -33.6236
22.7811 | -35.4237 | -33.6297
22.7872 | -35.4242 | -33.6359
22.7933 | -35.4246 | -33.6418
22.7994 | -35.4251 | -33.648
22.8055 | -35.4256 | -33.6542
22.8116 | -35.426 | -33.6605
22.8177 | -35.4265 | -33.6668
22.8238 | -35.4269 | -33.6729
22.8299 | -35.4274 | -33.6789
22.836 | -35.4279 | -33.685
22.8421 | -35.4283 | -33.6913
22.8482 | -35.4288 | -33.6975
22.8544 | -35.4293 | -33.7038
22.8605 | -35.4298 | -33.71
22.8666 | -35.4302 | -33.716
22.8727 | -35.4307 | -33.7223
22.8788 | -35.4311 | -33.7284
22.8849 | -35.4315 | -33.7346
22.891 | -35.432 | -33.7407
22.8971 | -35.4325 | -33.747
22.9032 | -35.4329 | -33.7529
22.9093 | -35.4334 | -33.7592
22.9154 | -35.4339 | -33.7652
22.9215 | -35.4343 | -33.7714
22.9276 | -35.4347 | -33.7776
22.9337 | -35.4352 | -33.7834
22.9398 | -35.4357 | -33.7897
22.9459 | -35.4361 | -33.7959
22.952 | -35.4366 | -33.8019
22.9581 | -35.437 | -33.8082
22.9642 | -35.4375 | -33.8144
22.9703 | -35.4379 | -33.8204
22.9764 | -35.4384 | -33.8266
22.9825 | -35.4388 | -33.8329
22.9886 | -35.4393 | -33.839
22.9948 | -35.4398 | -33.8451
23.0009 | -35.4402 | -33.8513
23.007 | -35.4407 | -33.8574
23.0131 | -35.4412 | -33.8633
23.0192 | -35.4416 | -33.8694
23.0253 | -35.4421 | -33.8754
23.0314 | -35.4425 | -33.8816
23.0375 | -35.4429 | -33.8878
23.0436 | -35.4434 | -33.8938
23.0497 | -35.4438 | -33.8999
23.0558 | -35.4442 | -33.9062
23.0619 | -35.4447 | -33.9123
23.068 | -35.4452 | -33.9186
23.0741 | -35.4457 | -33.9244
23.0802 | -35.4461 | -33.9306
23.0863 | -35.4466 | -33.9369
23.0924 | -35.447 | -33.9432
23.0985 | -35.4475 | -33.9493
23.1046 | -35.448 | -33.9555
23.1107 | -35.4484 | -33.9617
23.1168 | -35.4489 | -33.9679
23.1229 | -35.4494 | -33.9741
23.129 | -35.4498 | -33.9802
23.1351 | -35.4503 | -33.9863
23.1413 | -35.4507 | -33.9923
23.1474 | -35.4512 | -33.9982
23.1535 | -35.4516 | -34.0042
23.1596 | -35.4521 | -34.0102
23.1657 | -35.4525 | -34.0165
23.1718 | -35.453 | -34.0226
23.1779 | -35.4534 | -34.0289
23.184 | -35.4539 | -34.0349
23.1901 | -35.4543 | -34.0408
23.1962 | -35.4548 | -34.047
23.2023 | -35.4552 | -34.0531
23.2084 | -35.4557 | -34.0591
23.2145 | -35.4561 | -34.0649
23.2206 | -35.4566 | -34.071
23.2267 | -35.457 | -34.0769
23.2328 | -35.4574 | -34.0831
23.2389 | -35.4579 | -34.0889
23.245 | -35.4583 | -34.0951
23.2511 | -35.4588 | -34.1009
23.2572 | -35.4592 | -34.1069
23.2633 | -35.4597 | -34.1132
23.2694 | -35.4601 | -34.1193
23.2755 | -35.4606 | -34.1254
23.2817 | -35.461 | -34.1315
23.2878 | -35.4615 | -34.1378
23.2939 | -35.462 | -34.1439
23.3 | -35.4625 | -34.15
23.3061 | -35.4629 | -34.156
23.3122 | -35.4634 | -34.1622
23.3183 | -35.4639 | -34.1682
23.3244 | -35.4643 | -34.1743
23.3305 | -35.4648 | -34.1805
23.3366 | -35.4652 | -34.1865
23.3427 | -35.4658 | -34.1925
23.3488 | -35.4662 | -34.1988
23.3549 | -35.4667 | -34.205
23.361 | -35.4671 | -34.2108
23.3671 | -35.4675 | -34.217
23.3732 | -35.468 | -34.2231
23.3793 | -35.4685 | -34.2291
23.3854 | -35.4689 | -34.2352
23.3915 | -35.4694 | -34.2411
23.3976 | -35.4698 | -34.2471
23.4037 | -35.4703 | -34.2531
23.4098 | -35.4707 | -34.2592
23.4159 | -35.4712 | -34.2653
23.422 | -35.4717 | -34.2715
23.4282 | -35.4722 | -34.2774
23.4343 | -35.4726 | -34.2833
23.4404 | -35.4731 | -34.2897
23.4465 | -35.4736 | -34.2957
23.4526 | -35.474 | -34.3018
23.4587 | -35.4745 | -34.3077
23.4648 | -35.4749 | -34.3135
23.4709 | -35.4753 | -34.3198
23.477 | -35.4758 | -34.3258
23.4831 | -35.4762 | -34.3319
23.4892 | -35.4767 | -34.3381
23.4953 | -35.4771 | -34.3441
23.5014 | -35.4776 | -34.3503
23.5075 | -35.478 | -34.3563
23.5136 | -35.4785 | -34.3622
23.5197 | -35.4789 | -34.3683
23.5258 | -35.4793 | -34.3743
23.5319 | -35.4798 | -34.3803
23.538 | -35.4802 | -34.3864
23.5441 | -35.4806 | -34.3926
23.5502 | -35.4811 | -34.3987
23.5563 | -35.4815 | -34.405
23.5624 | -35.482 | -34.4111
23.5686 | -35.4825 | -34.4169
23.5747 | -35.4829 | -34.4229
23.5808 | -35.4834 | -34.4288
23.5869 | -35.4838 | -34.4348
23.593 | -35.4843 | -34.4409
23.5991 | -35.4847 | -34.4472
23.6052 | -35.4852 | -34.4533
23.6113 | -35.4856 | -34.4595
23.6174 | -35.4861 | -34.4655
23.6235 | -35.4865 | -34.4716
23.6296 | -35.487 | -34.4774
23.6357 | -35.4874 | -34.4836
23.6418 | -35.4879 | -34.4898
23.6479 | -35.4883 | -34.4956
23.654 | -35.4887 | -34.5015
23.6601 | -35.4892 | -34.5075
23.6662 | -35.4896 | -34.5135
23.6723 | -35.4901 | -34.5194
23.6784 | -35.4906 | -34.5254
23.6845 | -35.491 | -34.5314
23.6906 | -35.4915 | -34.5372
23.6967 | -35.4919 | -34.5431
23.7028 | -35.4924 | -34.5492
23.7089 | -35.4928 | -34.555
23.7151 | -35.4933 | -34.561
23.7212 | -35.4937 | -34.5671
23.7273 | -35.4942 | -34.5731
23.7334 | -35.4946 | -34.579
23.7395 | -35.495 | -34.5854
23.7456 | -35.4955 | -34.5909
23.7517 | -35.4959 | -34.5971
23.7578 | -35.4964 | -34.6031
23.7639 | -35.4968 | -34.6092
23.77 | -35.4973 | -34.6153
23.7761 | -35.4977 | -34.6214
23.7822 | -35.4982 | -34.6273
23.7883 | -35.4986 | -34.6332
23.7944 | -35.4991 | -34.6393
23.8005 | -35.4995 | -34.6453
23.8066 | -35.4999 | -34.6511
23.8127 | -35.5003 | -34.657
23.8188 | -35.5008 | -34.6631
23.8249 | -35.5012 | -34.6691
23.831 | -35.5017 | -34.6749
23.8371 | -35.5021 | -34.6809
23.8432 | -35.5026 | -34.6869
23.8493 | -35.503 | -34.6928
23.8555 | -35.5034 | -34.6988
23.8616 | -35.5039 | -34.7047
23.8677 | -35.5043 | -34.7109
23.8738 | -35.5047 | -34.7168
23.8799 | -35.5052 | -34.7231
23.886 | -35.5056 | -34.7289
23.8921 | -35.506 | -34.7351
23.8982 | -35.5065 | -34.7408
23.9043 | -35.5069 | -34.7468
23.9104 | -35.5074 | -34.7527
23.9165 | -35.5078 | -34.7587
23.9226 | -35.5083 | -34.7642
23.9287 | -35.5087 | -34.7703
23.9348 | -35.5091 | -34.7763
23.9409 | -35.5096 | -34.7824
23.947 | -35.51 | -34.7884
23.9531 | -35.5105 | -34.794
23.9592 | -35.5109 | -34.8001
23.9653 | -35.5114 | -34.8061
23.9714 | -35.5118 | -34.812
23.9775 | -35.5122 | -34.818
23.9836 | -35.5126 | -34.8241
23.9897 | -35.5131 | -34.8301
23.9958 | -35.5135 | -34.8362
24.002 | -35.5139 | -34.8421
24.0081 | -35.5144 | -34.848
24.0142 | -35.5148 | -34.8541
24.0203 | -35.5153 | -34.8603
24.0264 | -35.5157 | -34.8662
24.0325 | -35.5162 | -34.8723
24.0386 | -35.5166 | -34.8783
24.0447 | -35.517 | -34.8841
24.0508 | -35.5175 | -34.8901
24.0569 | -35.5179 | -34.8962
24.063 | -35.5183 | -34.902
24.0691 | -35.5188 | -34.9079
24.0752 | -35.5192 | -34.914
24.0813 | -35.5196 | -34.9198
24.0874 | -35.5201 | -34.9258
24.0935 | -35.5205 | -34.9318
24.0996 | -35.5209 | -34.9376
24.1057 | -35.5213 | -34.9435
24.1118 | -35.5218 | -34.9494
24.1179 | -35.5222 | -34.9555
24.124 | -35.5226 | -34.9614
24.1301 | -35.523 | -34.9671
24.1362 | -35.5235 | -34.9729
24.1424 | -35.5239 | -34.9789
24.1485 | -35.5243 | -34.985
24.1546 | -35.5248 | -34.991
24.1607 | -35.5252 | -34.9971
24.1668 | -35.5256 | -35.0028
24.1729 | -35.5261 | -35.009
24.179 | -35.5265 | -35.0151
24.1851 | -35.5269 | -35.021
24.1912 | -35.5273 | -35.027
24.1973 | -35.5278 | -35.0331
24.2034 | -35.5282 | -35.0387
24.2095 | -35.5286 | -35.0447
24.2156 | -35.5291 | -35.0509
24.2217 | -35.5295 | -35.0569
24.2278 | -35.53 | -35.0629
24.2339 | -35.5304 | -35.0687
24.24 | -35.5308 | -35.0744
24.2461 | -35.5313 | -35.0802
24.2522 | -35.5317 | -35.0863
24.2583 | -35.5321 | -35.0921
24.2644 | -35.5326 | -35.0981
24.2705 | -35.533 | -35.1039
24.2766 | -35.5334 | -35.1097
24.2827 | -35.5339 | -35.1156
24.2889 | -35.5343 | -35.1215
24.295 | -35.5347 | -35.1275
24.3011 | -35.5352 | -35.1334
24.3072 | -35.5356 | -35.1391
24.3133 | -35.536 | -35.1451
24.3194 | -35.5365 | -35.1512
24.3255 | -35.5369 | -35.1571
24.3316 | -35.5373 | -35.163
24.3377 | -35.5378 | -35.1688
24.3438 | -35.5382 | -35.1745
24.3499 | -35.5387 | -35.1803
24.356 | -35.5391 | -35.1862
24.3621 | -35.5395 | -35.192
24.3682 | -35.54 | -35.1978
24.3743 | -35.5404 | -35.2035
24.3804 | -35.5409 | -35.2089
24.3865 | -35.5413 | -35.2151
24.3926 | -35.5417 | -35.2209
24.3987 | -35.5421 | -35.2269
24.4048 | -35.5426 | -35.2327
24.4109 | -35.543 | -35.2384
24.417 | -35.5434 | -35.2443
24.4231 | -35.5439 | -35.2504
24.4293 | -35.5443 | -35.2562
24.4354 | -35.5447 | -35.2621
24.4415 | -35.5452 | -35.2682
24.4476 | -35.5456 | -35.2738
24.4537 | -35.546 | -35.2795
24.4598 | -35.5465 | -35.2856
24.4659 | -35.5469 | -35.2916
24.472 | -35.5473 | -35.2974
24.4781 | -35.5478 | -35.3032
24.4842 | -35.5482 | -35.309
24.4903 | -35.5486 | -35.3149
24.4964 | -35.549 | -35.3206
24.5025 | -35.5495 | -35.3265
24.5086 | -35.5499 | -35.3323
24.5147 | -35.5504 | -35.3381
24.5208 | -35.5508 | -35.3437
24.5269 | -35.5512 | -35.3495
24.533 | -35.5517 | -35.3555
24.5391 | -35.5521 | -35.3616
24.5452 | -35.5525 | -35.3673
24.5513 | -35.5529 | -35.3729
24.5574 | -35.5534 | -35.3787
24.5635 | -35.5538 | -35.3846
24.5696 | -35.5542 | -35.3905
24.5758 | -35.5547 | -35.3964
24.5819 | -35.5551 | -35.4021
24.588 | -35.5555 | -35.4079
24.5941 | -35.5559 | -35.4138
24.6002 | -35.5564 | -35.4196
24.6063 | -35.5568 | -35.4256
24.6124 | -35.5572 | -35.4316
24.6185 | -35.5577 | -35.4372
24.6246 | -35.5582 | -35.4431
24.6307 | -35.5586 | -35.4487
24.6368 | -35.5591 | -35.4546
24.6429 | -35.5595 | -35.4605
24.649 | -35.5599 | -35.4664
24.6551 | -35.5604 | -35.4724
24.6612 | -35.5609 | -35.4783
24.6673 | -35.5613 | -35.484
24.6734 | -35.5617 | -35.4898
24.6795 | -35.5621 | -35.4958
24.6856 | -35.5625 | -35.5015
24.6917 | -35.563 | -35.5074
24.6978 | -35.5634 | -35.5132
24.7039 | -35.5638 | -35.5191
24.71 | -35.5642 | -35.5249
24.7162 | -35.5646 | -35.5307
24.7223 | -35.5651 | -35.5366
24.7284 | -35.5655 | -35.5424
24.7345 | -35.5659 | -35.5483
24.7406 | -35.5663 | -35.5544
24.7467 | -35.5668 | -35.5604
24.7528 | -35.5672 | -35.5663
24.7589 | -35.5677 | -35.5722
24.765 | -35.5681 | -35.5776
24.7711 | -35.5685 | -35.5834
24.7772 | -35.5689 | -35.5892
24.7833 | -35.5694 | -35.5949
24.7894 | -35.5698 | -35.6008
24.7955 | -35.5702 | -35.6066
24.8016 | -35.5706 | -35.6124
24.8077 | -35.5711 | -35.6183
24.8138 | -35.5715 | -35.6245
24.8199 | -35.572 | -35.6303
24.826 | -35.5724 | -35.6361
24.8321 | -35.5728 | -35.6419
24.8382 | -35.5733 | -35.6475
24.8443 | -35.5737 | -35.6533
24.8504 | -35.5741 | -35.659
24.8565 | -35.5746 | -35.665
24.8627 | -35.575 | -35.6708
24.8688 | -35.5754 | -35.6765
24.8749 | -35.5759 | -35.6823
24.881 | -35.5763 | -35.6882
24.8871 | -35.5767 | -35.6942
24.8932 | -35.5772 | -35.7002
24.8993 | -35.5776 | -35.706
24.9054 | -35.578 | -35.7117
24.9115 | -35.5785 | -35.7175
24.9176 | -35.5789 | -35.7234
24.9237 | -35.5793 | -35.7292
24.9298 | -35.5797 | -35.7351
24.9359 | -35.5802 | -35.741
24.942 | -35.5806 | -35.7466
24.9481 | -35.5809 | -35.7525
24.9542 | -35.5814 | -35.7582
24.9603 | -35.5818 | -35.7641
24.9664 | -35.5822 | -35.7698
24.9725 | -35.5826 | -35.7757
24.9786 | -35.5831 | -35.7811
24.9847 | -35.5835 | -35.787
24.9908 | -35.5839 | -35.7927
24.9969 | -35.5843 | -35.7985
25.0031 | -35.5847 | -35.8044
25.0092 | -35.5851 | -35.81
25.0153 | -35.5856 | -35.8156
25.0214 | -35.586 | -35.8214
25.0275 | -35.5864 | -35.8273
25.0336 | -35.5868 | -35.8332
25.0397 | -35.5873 | -35.8392
25.0458 | -35.5877 | -35.8451
25.0519 | -35.5881 | -35.8507
25.058 | -35.5886 | -35.8565
25.0641 | -35.589 | -35.8622
25.0702 | -35.5894 | -35.8677
25.0763 | -35.5899 | -35.8734
25.0824 | -35.5902 | -35.879
25.0885 | -35.5907 | -35.8846
25.0946 | -35.5911 | -35.8902
25.1007 | -35.5915 | -35.8962
25.1068 | -35.5919 | -35.902
25.1129 | -35.5923 | -35.9077
25.119 | -35.5927 | -35.9133
25.1251 | -35.5932 | -35.919
25.1312 | -35.5936 | -35.9247
25.1373 | -35.594 | -35.9304
25.1435 | -35.5945 | -35.9362
25.1496 | -35.5949 | -35.9418
25.1557 | -35.5953 | -35.9475
25.1618 | -35.5958 | -35.9533
25.1679 | -35.5961 | -35.959
25.174 | -35.5966 | -35.9647
25.1801 | -35.597 | -35.9706
25.1862 | -35.5974 | -35.9763
25.1923 | -35.5978 | -35.9823
25.1984 | -35.5982 | -35.9881
25.2045 | -35.5986 | -35.994
25.2106 | -35.599 | -36
25.2167 | -35.5994 | -36.0057
25.2228 | -35.5998 | -36.0114
25.2289 | -35.6002 | -36.0172
25.235 | -35.6006 | -36.0228
25.2411 | -35.6011 | -36.0287
25.2472 | -35.6015 | -36.0345
25.2533 | -35.6019 | -36.0403
25.2594 | -35.6023 | -36.0461
25.2655 | -35.6027 | -36.0517
25.2716 | -35.6032 | -36.0574
25.2777 | -35.6036 | -36.0632
25.2838 | -35.604 | -36.069
25.29 | -35.6044 | -36.0747
25.2961 | -35.6048 | -36.0803
25.3022 | -35.6053 | -36.0861
25.3083 | -35.6057 | -36.0916
25.3144 | -35.6061 | -36.0973
25.3205 | -35.6065 | -36.1032
25.3266 | -35.6069 | -36.1089
25.3327 | -35.6074 | -36.1147
25.3388 | -35.6078 | -36.1204
25.3449 | -35.6083 | -36.1261
25.351 | -35.6087 | -36.132
25.3571 | -35.6091 | -36.1378
25.3632 | -35.6095 | -36.1432
25.3693 | -35.6099 | -36.149
25.3754 | -35.6103 | -36.1546
25.3815 | -35.6107 | -36.1604
25.3876 | -35.6111 | -36.1662
25.3937 | -35.6115 | -36.1719
25.3998 | -35.6119 | -36.1772
25.4059 | -35.6123 | -36.1831
25.412 | -35.6127 | -36.1886
25.4181 | -35.6131 | -36.1942
25.4242 | -35.6135 | -36.2
25.4304 | -35.614 | -36.2058
25.4365 | -35.6143 | -36.2113
25.4426 | -35.6147 | -36.217
25.4487 | -35.6151 | -36.2225
25.4548 | -35.6155 | -36.2282
25.4609 | -35.6159 | -36.2341
25.467 | -35.6163 | -36.2396
25.4731 | -35.6167 | -36.2454
25.4792 | -35.6171 | -36.2512
25.4853 | -35.6175 | -36.2569
25.4914 | -35.6179 | -36.2625
25.4975 | -35.6184 | -36.268
25.5036 | -35.6188 | -36.2737
25.5097 | -35.6192 | -36.2792
25.5158 | -35.6196 | -36.2849
25.5219 | -35.62 | -36.2906
25.528 | -35.6204 | -36.2963
25.5341 | -35.6208 | -36.302
25.5402 | -35.6213 | -36.3071
25.5463 | -35.6217 | -36.3127
25.5524 | -35.6221 | -36.3183
25.5585 | -35.6225 | -36.3238
25.5646 | -35.6229 | -36.3295
25.5707 | -35.6233 | -36.3352
25.5769 | -35.6237 | -36.3408
25.583 | -35.6241 | -36.3467
25.5891 | -35.6245 | -36.3523
25.5952 | -35.6249 | -36.358
25.6013 | -35.6253 | -36.3636
25.6074 | -35.6257 | -36.369
25.6135 | -35.626 | -36.3748
25.6196 | -35.6265 | -36.3804
25.6257 | -35.6269 | -36.3862
25.6318 | -35.6273 | -36.3916
25.6379 | -35.6277 | -36.3974
25.644 | -35.6281 | -36.4029
25.6501 | -35.6285 | -36.4085
25.6562 | -35.6289 | -36.4143
25.6623 | -35.6294 | -36.4199
25.6684 | -35.6298 | -36.4257
25.6745 | -35.6302 | -36.4315
25.6806 | -35.6306 | -36.4368
25.6867 | -35.631 | -36.4424
25.6928 | -35.6314 | -36.4481
25.6989 | -35.6318 | -36.4539
25.705 | -35.6322 | -36.4597
25.7111 | -35.6326 | -36.4652
25.7173 | -35.633 | -36.4706
25.7234 | -35.6335 | -36.4762
25.7295 | -35.6338 | -36.4819
25.7356 | -35.6342 | -36.4875
25.7417 | -35.6347 | -36.4932
25.7478 | -35.635 | -36.4987
25.7539 | -35.6355 | -36.5044
25.76 | -35.6358 | -36.5098
25.7661 | -35.6363 | -36.5153
25.7722 | -35.6367 | -36.5211
25.7783 | -35.6371 | -36.5271
25.7844 | -35.6375 | -36.5325
25.7905 | -35.6379 | -36.5381
25.7966 | -35.6383 | -36.5436
25.8027 | -35.6387 | -36.5492
25.8088 | -35.6391 | -36.5549
25.8149 | -35.6395 | -36.5605
25.821 | -35.6399 | -36.566
25.8271 | -35.6404 | -36.5716
25.8332 | -35.6408 | -36.5773
25.8393 | -35.6412 | -36.5832
25.8454 | -35.6416 | -36.5889
25.8515 | -35.642 | -36.5946
25.8576 | -35.6424 | -36.6002
25.8638 | -35.6428 | -36.6058
25.8699 | -35.6432 | -36.6114
25.876 | -35.6436 | -36.6171
25.8821 | -35.644 | -36.6228
25.8882 | -35.6443 | -36.6282
25.8943 | -35.6448 | -36.6338
25.9004 | -35.6452 | -36.6394
25.9065 | -35.6456 | -36.6451
25.9126 | -35.646 | -36.6508
25.9187 | -35.6464 | -36.6564
25.9248 | -35.6468 | -36.662
25.9309 | -35.6471 | -36.6677
25.937 | -35.6476 | -36.6734
25.9431 | -35.648 | -36.679
25.9492 | -35.6484 | -36.6846
25.9553 | -35.6488 | -36.6904
25.9614 | -35.6492 | -36.6957
25.9675 | -35.6496 | -36.7014
25.9736 | -35.65 | -36.7071
25.9797 | -35.6503 | -36.7126
25.9858 | -35.6508 | -36.7181
25.9919 | -35.6512 | -36.7235
25.998 | -35.6515 | -36.7289
26.0042 | -35.6519 | -36.7345
26.0103 | -35.6524 | -36.7401
26.0164 | -35.6527 | -36.7459
26.0225 | -35.6531 | -36.7514
26.0286 | -35.6535 | -36.7569
26.0347 | -35.6539 | -36.7623
26.0408 | -35.6543 | -36.7678
26.0469 | -35.6547 | -36.7732
26.053 | -35.6551 | -36.7789
26.0591 | -35.6554 | -36.7845
26.0652 | -35.6558 | -36.79
26.0713 | -35.6562 | -36.7955
26.0774 | -35.6566 | -36.8011
26.0835 | -35.657 | -36.8068
26.0896 | -35.6574 | -36.8123
26.0957 | -35.6578 | -36.8178
26.1018 | -35.6582 | -36.8234
26.1079 | -35.6586 | -36.8289
26.114 | -35.659 | -36.8344
26.1201 | -35.6593 | -36.8401
26.1262 | -35.6598 | -36.8457
26.1323 | -35.6602 | -36.8514
26.1384 | -35.6605 | -36.8565
26.1445 | -35.6609 | -36.8621
26.1507 | -35.6614 | -36.8676
26.1568 | -35.6617 | -36.8735
26.1629 | -35.6621 | -36.8789
26.169 | -35.6625 | -36.8844
26.1751 | -35.6629 | -36.8899
26.1812 | -35.6633 | -36.8953
26.1873 | -35.6637 | -36.9008
26.1934 | -35.6642 | -36.9064
26.1995 | -35.6646 | -36.912
26.2056 | -35.6649 | -36.9172
26.2117 | -35.6653 | -36.9226
26.2178 | -35.6658 | -36.9281
26.2239 | -35.6662 | -36.9336
26.23 | -35.6666 | -36.9393
26.2361 | -35.6669 | -36.9449
26.2422 | -35.6673 | -36.9504
26.2483 | -35.6677 | -36.9558
26.2544 | -35.6681 | -36.9611
26.2605 | -35.6685 | -36.9666
26.2666 | -35.6689 | -36.9722
26.2727 | -35.6693 | -36.9778
26.2788 | -35.6697 | -36.9832
26.2849 | -35.6701 | -36.9887
26.2911 | -35.6705 | -36.9943
26.2972 | -35.6709 | -36.9999
26.3033 | -35.6713 | -37.0054
26.3094 | -35.6716 | -37.011
26.3155 | -35.672 | -37.0164
26.3216 | -35.6724 | -37.0221
26.3277 | -35.6728 | -37.0274
26.3338 | -35.6732 | -37.0329
26.3399 | -35.6735 | -37.0385
26.346 | -35.6739 | -37.0435
26.3521 | -35.6743 | -37.0489
26.3582 | -35.6746 | -37.0546
26.3643 | -35.675 | -37.0602
26.3704 | -35.6755 | -37.0657
26.3765 | -35.6758 | -37.0712
26.3826 | -35.6762 | -37.0766
26.3887 | -35.6766 | -37.0822
26.3948 | -35.677 | -37.0879
26.4009 | -35.6773 | -37.0935
26.407 | -35.6777 | -37.099
26.4131 | -35.6781 | -37.1046
26.4192 | -35.6785 | -37.1099
26.4253 | -35.6789 | -37.1156
26.4314 | -35.6793 | -37.1211
26.4376 | -35.6797 | -37.1267
26.4437 | -35.6801 | -37.1323
26.4498 | -35.6804 | -37.138
26.4559 | -35.6808 | -37.1432
26.462 | -35.6812 | -37.1487
26.4681 | -35.6816 | -37.1543
26.4742 | -35.682 | -37.16
26.4803 | -35.6823 | -37.1655
26.4864 | -35.6827 | -37.1708
26.4925 | -35.6831 | -37.1762
26.4986 | -35.6835 | -37.1816
26.5047 | -35.6839 | -37.1871
26.5108 | -35.6843 | -37.1927
26.5169 | -35.6847 | -37.1981
26.523 | -35.6851 | -37.2037
26.5291 | -35.6855 | -37.2092
26.5352 | -35.6858 | -37.2148
26.5413 | -35.6862 | -37.2205
26.5474 | -35.6866 | -37.2261
26.5535 | -35.687 | -37.2315
26.5596 | -35.6873 | -37.2369
26.5657 | -35.6877 | -37.2423
26.5718 | -35.6881 | -37.2477
26.578 | -35.6885 | -37.2531
26.5841 | -35.6889 | -37.2585
26.5902 | -35.6893 | -37.264
26.5963 | -35.6896 | -37.2695
26.6024 | -35.69 | -37.275
26.6085 | -35.6904 | -37.2808
26.6146 | -35.6908 | -37.2865
26.6207 | -35.6912 | -37.2919
26.6268 | -35.6916 | -37.2975
26.6329 | -35.6919 | -37.3029
26.639 | -35.6923 | -37.3085
26.6451 | -35.6927 | -37.3142
26.6512 | -35.6931 | -37.3196
26.6573 | -35.6934 | -37.3249
26.6634 | -35.6938 | -37.3299
26.6695 | -35.6941 | -37.3355
26.6756 | -35.6945 | -37.341
26.6817 | -35.6949 | -37.3464
26.6878 | -35.6953 | -37.3519
26.6939 | -35.6957 | -37.3574
26.7 | -35.6961 | -37.3628
26.7061 | -35.6964 | -37.3682
26.7122 | -35.6968 | -37.3738
26.7183 | -35.6972 | -37.3793
26.7245 | -35.6975 | -37.3847
26.7306 | -35.6979 | -37.3902
26.7367 | -35.6983 | -37.3954
26.7428 | -35.6987 | -37.4009
26.7489 | -35.699 | -37.4062
26.755 | -35.6994 | -37.4117
26.7611 | -35.6998 | -37.4173
26.7672 | -35.7002 | -37.4228
26.7733 | -35.7005 | -37.4281
26.7794 | -35.7009 | -37.4336
26.7855 | -35.7013 | -37.4391
26.7916 | -35.7017 | -37.4443
26.7977 | -35.7021 | -37.4497
26.8038 | -35.7024 | -37.4552
26.8099 | -35.7028 | -37.4608
26.816 | -35.7032 | -37.4662
26.8221 | -35.7035 | -37.4717
26.8282 | -35.7039 | -37.4772
26.8343 | -35.7043 | -37.4827
26.8404 | -35.7047 | -37.4881
26.8465 | -35.7051 | -37.4936
26.8526 | -35.7055 | -37.4991
26.8587 | -35.7059 | -37.5046
26.8649 | -35.7062 | -37.5099
26.871 | -35.7066 | -37.5156
26.8771 | -35.707 | -37.521
26.8832 | -35.7074 | -37.5264
26.8893 | -35.7077 | -37.5317
26.8954 | -35.7081 | -37.537
26.9015 | -35.7085 | -37.5425
26.9076 | -35.7089 | -37.548
26.9137 | -35.7092 | -37.5535
26.9198 | -35.7096 | -37.5593
26.9259 | -35.71 | -37.5649
26.932 | -35.7103 | -37.5705
26.9381 | -35.7107 | -37.576
26.9442 | -35.7111 | -37.5812
26.9503 | -35.7114 | -37.5866
26.9564 | -35.7118 | -37.5922
26.9625 | -35.7121 | -37.5977
26.9686 | -35.7125 | -37.6032
26.9747 | -35.7129 | -37.6084
26.9808 | -35.7133 | -37.6136
26.9869 | -35.7136 | -37.619
26.993 | -35.714 | -37.6246
26.9991 | -35.7144 | -37.6302
27.0052 | -35.7148 | -37.6355
27.0114 | -35.7152 | -37.6408
27.0175 | -35.7155 | -37.646
27.0236 | -35.7159 | -37.6513
27.0297 | -35.7162 | -37.6567
27.0358 | -35.7166 | -37.6622
27.0419 | -35.717 | -37.6674
27.048 | -35.7173 | -37.6729
27.0541 | -35.7177 | -37.6781
27.0602 | -35.7181 | -37.6838
27.0663 | -35.7184 | -37.6891
27.0724 | -35.7188 | -37.6945
27.0785 | -35.7192 | -37.7001
27.0846 | -35.7195 | -37.7052
27.0907 | -35.7199 | -37.7107
27.0968 | -35.7203 | -37.7162
27.1029 | -35.7206 | -37.7216
27.109 | -35.721 | -37.727
27.1151 | -35.7214 | -37.7325
27.1212 | -35.7217 | -37.7376
27.1273 | -35.7221 | -37.7432
27.1334 | -35.7225 | -37.7489
27.1395 | -35.7229 | -37.7544
27.1456 | -35.7232 | -37.7595
27.1518 | -35.7236 | -37.7649
27.1579 | -35.7239 | -37.7701
27.164 | -35.7243 | -37.7753
27.1701 | -35.7246 | -37.7807
27.1762 | -35.725 | -37.7861
27.1823 | -35.7253 | -37.7916
27.1884 | -35.7257 | -37.7971
27.1945 | -35.726 | -37.8023
27.2006 | -35.7264 | -37.8077
27.2067 | -35.7267 | -37.8129
27.2128 | -35.7271 | -37.8184
27.2189 | -35.7275 | -37.8238
27.225 | -35.7278 | -37.8291
27.2311 | -35.7282 | -37.8346
27.2372 | -35.7286 | -37.8401
27.2433 | -35.729 | -37.8454
27.2494 | -35.7293 | -37.8509
27.2555 | -35.7297 | -37.8563
27.2616 | -35.73 | -37.8615
27.2677 | -35.7304 | -37.867
27.2738 | -35.7308 | -37.8723
27.2799 | -35.7311 | -37.8776
27.286 | -35.7315 | -37.8828
27.2921 | -35.7318 | -37.8883
27.2983 | -35.7322 | -37.8938
27.3044 | -35.7326 | -37.8991
27.3105 | -35.7329 | -37.9042
27.3166 | -35.7333 | -37.9095
27.3227 | -35.7337 | -37.9151
27.3288 | -35.7341 | -37.9206
27.3349 | -35.7344 | -37.9257
27.341 | -35.7347 | -37.931
27.3471 | -35.7351 | -37.9363
27.3532 | -35.7354 | -37.9414
27.3593 | -35.7358 | -37.9469
27.3654 | -35.7362 | -37.9523
27.3715 | -35.7365 | -37.9574
27.3776 | -35.7369 | -37.9628
27.3837 | -35.7373 | -37.9681
27.3898 | -35.7377 | -37.9734
27.3959 | -35.738 | -37.9788
27.402 | -35.7385 | -37.9838
27.4081 | -35.7388 | -37.9892
27.4142 | -35.7392 | -37.9947
27.4203 | -35.7395 | -38.0002
27.4264 | -35.7399 | -38.0056
27.4325 | -35.7403 | -38.0108
27.4387 | -35.7406 | -38.0164
27.4448 | -35.7409 | -38.0217
27.4509 | -35.7413 | -38.0271
27.457 | -35.7417 | -38.0323
27.4631 | -35.742 | -38.0378
27.4692 | -35.7424 | -38.0431
27.4753 | -35.7427 | -38.0484
27.4814 | -35.743 | -38.0537
27.4875 | -35.7434 | -38.0588
27.4936 | -35.7437 | -38.0643
27.4997 | -35.7441 | -38.0696
27.5058 | -35.7444 | -38.0751
27.5119 | -35.7447 | -38.0801
27.518 | -35.745 | -38.0852
27.5241 | -35.7454 | -38.0908
27.5302 | -35.7458 | -38.0962
27.5363 | -35.7461 | -38.1015
27.5424 | -35.7465 | -38.1068
27.5485 | -35.7469 | -38.1119
27.5546 | -35.7472 | -38.1174
27.5607 | -35.7476 | -38.1227
27.5668 | -35.7479 | -38.128
27.5729 | -35.7483 | -38.1332
27.5791 | -35.7486 | -38.1384
27.5852 | -35.7489 | -38.1439
27.5913 | -35.7493 | -38.1491
27.5974 | -35.7496 | -38.1544
27.6035 | -35.75 | -38.1598
27.6096 | -35.7503 | -38.165
27.6157 | -35.7506 | -38.1704
27.6218 | -35.751 | -38.1756
27.6279 | -35.7514 | -38.1809
27.634 | -35.7517 | -38.1863
27.6401 | -35.7521 | -38.1916
27.6462 | -35.7524 | -38.197
27.6523 | -35.7528 | -38.2019
27.6584 | -35.7531 | -38.2073
27.6645 | -35.7534 | -38.2122
27.6706 | -35.7538 | -38.2176
27.6767 | -35.7542 | -38.2233
27.6828 | -35.7545 | -38.2284
27.6889 | -35.7548 | -38.2336
27.695 | -35.7552 | -38.2389
27.7011 | -35.7555 | -38.2443
27.7072 | -35.7558 | -38.2499
27.7133 | -35.7562 | -38.2551
27.7194 | -35.7565 | -38.2603
27.7256 | -35.7569 | -38.2658
27.7317 | -35.7572 | -38.271
27.7378 | -35.7576 | -38.2763
27.7439 | -35.758 | -38.2817
27.75 | -35.7584 | -38.2871
27.7561 | -35.7587 | -38.2924
27.7622 | -35.7591 | -38.2977
27.7683 | -35.7594 | -38.3031
27.7744 | -35.7598 | -38.3079
27.7805 | -35.7602 | -38.3132
27.7866 | -35.7606 | -38.3182
27.7927 | -35.7609 | -38.3228
27.7988 | -35.7613 | -38.3281
27.8049 | -35.7617 | -38.3334
27.811 | -35.7621 | -38.3386
27.8171 | -35.7624 | -38.3439
27.8232 | -35.7628 | -38.3491
27.8293 | -35.7632 | -38.3544
27.8354 | -35.7635 | -38.3599
27.8415 | -35.7639 | -38.3649
27.8476 | -35.7642 | -38.3706
27.8537 | -35.7646 | -38.376
27.8598 | -35.7649 | -38.3812
27.866 | -35.7653 | -38.3869
27.8721 | -35.7656 | -38.3925
27.8782 | -35.766 | -38.3977
27.8843 | -35.7663 | -38.4035
27.8904 | -35.7666 | -38.4089
27.8965 | -35.767 | -38.414
27.9026 | -35.7673 | -38.4195
27.9087 | -35.7676 | -38.4247
27.9148 | -35.768 | -38.4299
27.9209 | -35.7683 | -38.4352
27.927 | -35.7686 | -38.4403
27.9331 | -35.769 | -38.4454
27.9392 | -35.7693 | -38.4505
27.9453 | -35.7696 | -38.4558
27.9514 | -35.77 | -38.4609
27.9575 | -35.7703 | -38.4661
27.9636 | -35.7707 | -38.4713
27.9697 | -35.771 | -38.4767
27.9758 | -35.7714 | -38.4822
27.9819 | -35.7717 | -38.4871
27.988 | -35.772 | -38.4923
27.9941 | -35.7724 | -38.4975
28.0002 | -35.7727 | -38.5025
28.0063 | -35.7731 | -38.5078
28.0125 | -35.7735 | -38.5131
28.0186 | -35.7738 | -38.5182
28.0247 | -35.7741 | -38.5236
28.0308 | -35.7745 | -38.5288
28.0369 | -35.7748 | -38.5341
28.043 | -35.7752 | -38.5397
28.0491 | -35.7755 | -38.5449
28.0552 | -35.7759 | -38.5504
28.0613 | -35.7762 | -38.5556
28.0674 | -35.7766 | -38.561
28.0735 | -35.7769 | -38.5663

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5133 | -43.3823
21.4908 | -35.5191 | -35.3437
16.4948 | -34.8848 | -27.8978
12.6602 | -34.5105 | -21.5806
9.7171 | -34.2898 | -16.6367
7.4582 | -34.1309 | -12.6268
5.7244 | -34.0242 | -9.3823
4.3936 | -33.9577 | -6.7618
3.3722 | -33.9411 | -4.7055
2.5883 | -33.9394 | -3.1437
1.9866 | -33.9369 | -1.8997
1.5248 | -33.9394 | -0.90354
1.1703 | -33.9482 | -0.25566
1 | -33.948 | 0.067894
0.89822 | -33.9294 | 0.33829
0.68941 | -33.9135 | 0.91372
0.52919 | -33.9208 | 1.3588
0.40612 | -33.914 | 1.9615
0.31171 | -33.9054 | 2.409
0.23928 | -33.9116 | 3.1257
0.18358 | -33.9189 | 3.5359
0.14091 | -33.9302 | 4.4009
0.10819 | -33.9432 | 5.5512
0.083017 | -33.9677 | 7.0141
0.063705 | -34.0077 | 8.8677
0.048923 | -34.0677 | 11.3321
0.037479 | -34.1749 | 14.5052
0.028801 | -34.3493 | 18.5234
0.022078 | -34.6541 | 23.9217
0.017023 | -35.1313 | 30.3511
0.013018 | -35.9358 | 38.0274
0.010014 | -37.0514 | 45.9397

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003751 | -12.1288
0.007501 | -29.3369
0.011252 | -34.6457
0.015002 | -38.3884
0.018753 | -41.3543
0.022504 | -42.7717
0.026254 | -43.7821
0.030005 | -45.7142
0.033755 | -46.1776
0.037506 | -47.9672
0.041256 | -49.6712
0.045007 | -50.723
0.048758 | -50.7382
0.052508 | -50.8824
0.056259 | -51.4031
0.060009 | -52.8072
0.06376 | -54.0746
0.067511 | -54.764
0.071261 | -55.8584
0.075012 | -55.9348
0.078762 | -56.5777
0.082513 | -56.7397
0.086263 | -58.1512
0.090014 | -59.6727
0.093765 | -60.135
0.097515 | -60.2833
0.10127 | -61.4519
0.10502 | -60.8586
0.10877 | -60.9888
0.11252 | -61.0777
0.11627 | -62.5904
0.12002 | -62.8962
0.12377 | -62.6999
0.12752 | -62.8473
0.13127 | -62.7171
0.13502 | -64.2585
0.13877 | -65.0263
0.14252 | -64.0559
0.14627 | -63.977
0.15002 | -65.2281
0.15377 | -65.6215
0.15752 | -64.5323
0.16128 | -65.4406
0.16503 | -67.0912
0.16878 | -67.2933
0.17253 | -67.5746
0.17628 | -67.7328
0.18003 | -66.3408
0.18378 | -66.6616
0.18753 | -67.6262
0.19128 | -68.0393
0.19503 | -68.8227
0.19878 | -70.1252
0.20253 | -69.7185
0.20628 | -68.8967
0.21003 | -70.0005
0.21378 | -70.4598
0.21753 | -70.9505
0.22129 | -71.3202
0.22504 | -71.5602
0.22879 | -71.6507
0.23254 | -71.3906
0.23629 | -72.1591
0.24004 | -71.8004
0.24379 | -70.9107
0.24754 | -71.5846
0.25129 | -72.1525
0.25504 | -71.9159
0.25879 | -71.5488
0.26254 | -71.9419
0.26629 | -73.1603
0.27004 | -73.8936
0.27379 | -73.075
0.27754 | -72.3983
0.28129 | -73.1498
0.28504 | -73.431
0.2888 | -73.5828
0.29255 | -73.3817
0.2963 | -74.9297
0.30005 | -74.7838
0.3038 | -73.4404
0.30755 | -73.5367
0.3113 | -74.2784
0.31505 | -74.4069
0.3188 | -73.9571
0.32255 | -73.8164
0.3263 | -73.9355
0.33005 | -74.4364
0.3338 | -74.7393
0.33755 | -75.2386
0.3413 | -75.4939
0.34505 | -76.1282
0.3488 | -76.6599
0.35256 | -76.1565
0.35631 | -75.4777
0.36006 | -76.2809
0.36381 | -76.6534
0.36756 | -76.6772
0.37131 | -76.5353
0.37506 | -76.0512
0.37881 | -75.9509
0.38256 | -76.392
0.38631 | -76.456
0.39006 | -76.9345
0.39381 | -77.3523
0.39756 | -77.8408
0.40131 | -78.7671
0.40506 | -78.0756
0.40881 | -78.6053
0.41256 | -77.2468
0.41631 | -77.4734
0.42007 | -77.6024
0.42382 | -76.0054
0.42757 | -75.6969
0.43132 | -76.9494
0.43507 | -78.1442
0.43882 | -78.6761
0.44257 | -78.4663
0.44632 | -78.5995
0.45007 | -79.378
0.45382 | -77.9472
0.45757 | -78.5395
0.46132 | -79.5802
0.46507 | -79.9484
0.46882 | -80.0143
0.47257 | -78.6656
0.47632 | -79.0358
0.48007 | -79.8304
0.48383 | -79.329
0.48758 | -78.6448
0.49133 | -78.437
0.49508 | -77.8504
0.49883 | -79.2877
0.50258 | -79.1944
0.50633 | -79.0599
0.51008 | -78.7046
0.51383 | -79.3787
0.51758 | -80.9095
0.52133 | -80.1733
0.52508 | -79.0286
0.52883 | -80.5182
0.53258 | -80.5686
0.53633 | -79.9223
0.54008 | -79.9096
0.54383 | -80.1519
0.54759 | -80.3277
0.55134 | -80.0453
0.55509 | -79.9052
0.55884 | -80.2536
0.56259 | -80.8126
0.56634 | -81.1617
0.57009 | -81.4093
0.57384 | -81.1452
0.57759 | -80.7922
0.58134 | -80.9711
0.58509 | -81.0313
0.58884 | -81.0133
0.59259 | -81.3875
0.59634 | -81.9444
0.60009 | -82.2549
0.60384 | -82.0261
0.60759 | -81.4305
0.61135 | -81.3169
0.6151 | -81.8408
0.61885 | -82.4214
0.6226 | -82.9825
0.62635 | -83.2965
0.6301 | -82.7621
0.63385 | -82.0982
0.6376 | -81.8623
0.64135 | -81.8613
0.6451 | -81.8283
0.64885 | -81.8543
0.6526 | -82.0003
0.65635 | -82.4371
0.6601 | -83.1394
0.66385 | -83.574
0.6676 | -83.218
0.67136 | -82.5942
0.6751 | -82.5262
0.67886 | -83.2614
0.68261 | -83.9263
0.68636 | -84.0559
0.69011 | -83.6025
0.69386 | -82.8283
0.69761 | -82.1253
0.70136 | -82.0585
0.70511 | -82.4045
0.70886 | -82.5312
0.71261 | -82.5988
0.71636 | -82.6482
0.72011 | -82.5763
0.72386 | -82.7024
0.72761 | -82.7922
0.73136 | -83.0514
0.73511 | -83.5386
0.73886 | -83.7206
0.74262 | -83.8175
0.74637 | -83.774
0.75012 | -83.8852
0.75387 | -83.9477
0.75762 | -83.7966
0.76137 | -83.6768
0.76512 | -83.8311
0.76887 | -84.1602
0.77262 | -84.4586
0.77637 | -84.1053
0.78012 | -83.7877
0.78387 | -83.9574
0.78762 | -84.1973
0.79137 | -84.6383
0.79512 | -84.934
0.79888 | -84.8545
0.80263 | -84.6266
0.80638 | -84.3631
0.81013 | -84.4239
0.81388 | -84.3257
0.81763 | -84.2306
0.82138 | -84.4636
0.82513 | -84.5128
0.82888 | -84.7913
0.83263 | -85.2929
0.83638 | -85.2235
0.84013 | -84.6916
0.84388 | -84.5893
0.84763 | -85.1932
0.85138 | -85.4892
0.85513 | -85.2963
0.85888 | -85.0013
0.86264 | -84.6694
0.86638 | -84.4599
0.87014 | -84.6321
0.87389 | -84.7676
0.87764 | -84.6578
0.88139 | -84.7419
0.88514 | -84.8217
0.88889 | -85.2667
0.89264 | -86.1001
0.89639 | -86.2899
0.90014 | -86.0239
0.90389 | -85.9175
0.90764 | -85.7428
0.91139 | -85.6284
0.91514 | -85.6946
0.91889 | -85.6043
0.92264 | -85.5564
0.92639 | -85.6065
0.93014 | -85.7499
0.9339 | -85.8723
0.93765 | -85.9885
0.9414 | -86.1679
0.94515 | -86.5282
0.9489 | -86.8361
0.95265 | -86.7981
0.9564 | -87.0189
0.96015 | -87.44
0.9639 | -87.2453
0.96765 | -86.5339
0.9714 | -85.9816
0.97515 | -85.7437
0.9789 | -85.8329
0.98265 | -86.1876
0.9864 | -86.5014
0.99016 | -86.3435
0.99391 | -86.1739
0.99766 | -86.5899
1.0014 | -87.1453
1.0052 | -87.1327
1.0089 | -86.7561
1.0127 | -86.7264
1.0164 | -87.1456
1.0202 | -87.321
1.0239 | -86.9576
1.0277 | -86.2573
1.0314 | -85.6285
1.0352 | -85.6528
1.0389 | -86.2419
1.0427 | -86.8888
1.0464 | -87.5482
1.0502 | -87.8288
1.0539 | -87.3805
1.0577 | -86.7813
1.0614 | -86.0487
1.0652 | -85.8037
1.0689 | -86.4186
1.0727 | -87.1178
1.0764 | -87.247
1.0802 | -87.1165
1.0839 | -87.1612
1.0877 | -87.4164
1.0914 | -87.8204
1.0952 | -88.1535
1.0989 | -88.3101
1.1027 | -88.3332
1.1064 | -88.2653
1.1102 | -88.059
1.1139 | -87.657
1.1177 | -87.2863
1.1214 | -87.1471
1.1252 | -87.2072
1.1289 | -87.3888
1.1327 | -87.5465
1.1364 | -87.5086
1.1402 | -87.3788
1.1439 | -87.4379
1.1477 | -87.6882
1.1514 | -87.8402
1.1552 | -87.8004
1.1589 | -87.7655
1.1627 | -87.8824
1.1664 | -88.0078
1.1702 | -87.9761
1.1739 | -88.0262
1.1777 | -88.3153
1.1814 | -88.4955
1.1852 | -88.2976
1.1889 | -87.9564
1.1927 | -87.7405
1.1964 | -87.6493
1.2002 | -87.5592
1.2039 | -87.5617
1.2077 | -87.8227
1.2114 | -88.1212
1.2152 | -88.1655
1.2189 | -88.1573
1.2227 | -88.3082
1.2264 | -88.5941
1.2302 | -88.8555
1.2339 | -88.8576
1.2377 | -88.6577
1.2414 | -88.4487
1.2452 | -88.2639
1.2489 | -88.2004
1.2527 | -88.3469
1.2564 | -88.6243
1.2602 | -88.9572
1.2639 | -89.2961
1.2677 | -89.4289
1.2714 | -89.2849
1.2752 | -89.091
1.279 | -88.9449
1.2827 | -88.7555
1.2865 | -88.5486
1.2902 | -88.5396
1.294 | -88.7995
1.2977 | -89.0748
1.3015 | -89.0775
1.3052 | -88.9047
1.309 | -88.801
1.3127 | -88.7828
1.3165 | -88.7336
1.3202 | -88.6691
1.324 | -88.66
1.3277 | -88.684
1.3315 | -88.7688
1.3352 | -88.972
1.339 | -89.1269
1.3427 | -89.0501
1.3465 | -88.8393
1.3502 | -88.7384
1.354 | -88.8558
1.3577 | -89.0986
1.3615 | -89.3093
1.3652 | -89.39
1.369 | -89.3087
1.3727 | -89.2552
1.3765 | -89.4332
1.3802 | -89.6837
1.384 | -89.782
1.3877 | -89.8402
1.3915 | -90.0083
1.3952 | -90.2863
1.399 | -90.4898
1.4027 | -90.4563
1.4065 | -90.3375
1.4102 | -90.3341
1.414 | -90.4324
1.4177 | -90.3634
1.4215 | -89.9685
1.4252 | -89.4781
1.429 | -89.1949
1.4327 | -89.2674
1.4365 | -89.6471
1.4402 | -89.9638
1.444 | -89.8576
1.4477 | -89.5882
1.4515 | -89.607
1.4552 | -89.9244
1.459 | -90.2899
1.4627 | -90.5627
1.4665 | -90.7059
1.4702 | -90.6405
1.474 | -90.464
1.4777 | -90.3736
1.4815 | -90.3366
1.4852 | -90.2123
1.489 | -89.9134
1.4927 | -89.4277
1.4965 | -89.0138
1.5002 | -88.9932
1.504 | -89.3076
1.5077 | -89.5328
1.5115 | -89.5102
1.5152 | -89.5445
1.519 | -89.8777
1.5227 | -90.4326
1.5265 | -90.8488
1.5302 | -90.8128
1.534 | -90.5223
1.5377 | -90.3715
1.5415 | -90.4504
1.5452 | -90.6098
1.549 | -90.7201
1.5527 | -90.7162
1.5565 | -90.6622
1.5602 | -90.5054
1.564 | -90.0346
1.5677 | -89.4253
1.5715 | -89.18
1.5752 | -89.5525
1.579 | -90.3378
1.5827 | -91.0268
1.5865 | -91.3121
1.5902 | -91.2476
1.594 | -90.9898
1.5978 | -90.5688
1.6015 | -90.0626
1.6053 | -89.7758
1.609 | -89.9103
1.6128 | -90.3708
1.6165 | -90.8641
1.6203 | -90.9848
1.624 | -91.0885
1.6278 | -91.0861
1.6315 | -91.0797
1.6353 | -91.1643
1.639 | -91.3492
1.6428 | -91.5316
1.6465 | -91.5906
1.6503 | -91.5113
1.654 | -91.3827
1.6578 | -91.2905
1.6615 | -91.2195
1.6653 | -91.1126
1.669 | -90.993
1.6728 | -90.9321
1.6765 | -90.9582
1.6803 | -91.065
1.684 | -91.2176
1.6878 | -91.3586
1.6915 | -91.4401
1.6953 | -91.4726
1.699 | -91.4361
1.7028 | -91.3106
1.7065 | -91.1763
1.7103 | -91.1151
1.714 | -91.1414
1.7178 | -91.2495
1.7215 | -91.3991
1.7253 | -91.5162
1.729 | -91.5231
1.7328 | -91.4874
1.7365 | -91.5366
1.7403 | -91.6947
1.744 | -91.8714
1.7478 | -91.9419
1.7515 | -91.9449
1.7553 | -92.0534
1.759 | -92.3027
1.7628 | -92.4769
1.7665 | -92.3232
1.7703 | -91.8975
1.774 | -91.4891
1.7778 | -91.3024
1.7815 | -91.3838
1.7853 | -91.643
1.789 | -91.9184
1.7928 | -92.0706
1.7965 | -92.1021
1.8003 | -92.1146
1.804 | -92.1733
1.8078 | -92.2325
1.8115 | -92.2699
1.8153 | -92.3162
1.819 | -92.4011
1.8228 | -92.4886
1.8265 | -92.5073
1.8303 | -92.4581
1.834 | -92.4093
1.8378 | -92.3927
1.8415 | -92.3872
1.8453 | -92.3501
1.849 | -92.2875
1.8528 | -92.2793
1.8565 | -92.3969
1.8603 | -92.6536
1.864 | -92.9507
1.8678 | -93.1524
1.8715 | -93.1612
1.8753 | -92.9964
1.879 | -92.7794
1.8828 | -92.6007
1.8865 | -92.4328
1.8903 | -92.26
1.894 | -92.1379
1.8978 | -92.1316
1.9015 | -92.188
1.9053 | -92.2113
1.909 | -92.1525
1.9128 | -92.054
1.9165 | -92.0215
1.9203 | -92.1404
1.9241 | -92.4104
1.9278 | -92.6543
1.9316 | -92.6345
1.9353 | -92.3465
1.9391 | -92.0049
1.9428 | -91.785
1.9466 | -91.7343
1.9503 | -91.8364
1.9541 | -92.0606
1.9578 | -92.3476
1.9616 | -92.658
1.9653 | -92.9264
1.9691 | -93.0841
1.9728 | -93.123
1.9766 | -93.1161
1.9803 | -93.12
1.9841 | -93.1199
1.9878 | -93.0828
1.9916 | -92.9761
1.9953 | -92.8273
1.9991 | -92.7028
2.0028 | -92.6572
2.0066 | -92.7469
2.0103 | -92.9571
2.0141 | -93.1371
2.0178 | -93.1237
2.0216 | -92.9276
2.0253 | -92.7547
2.0291 | -92.7566
2.0328 | -92.961
2.0366 | -93.2662
2.0403 | -93.488
2.0441 | -93.5346
2.0478 | -93.4588
2.0516 | -93.3711
2.0553 | -93.3519
2.0591 | -93.3772
2.0628 | -93.3518
2.0666 | -93.197
2.0703 | -92.9732
2.0741 | -92.7706
2.0778 | -92.6325
2.0816 | -92.5442
2.0853 | -92.4766
2.0891 | -92.4465
2.0928 | -92.4937
2.0966 | -92.6077
2.1003 | -92.7235
2.1041 | -92.7673
2.1078 | -92.7616
2.1116 | -92.7513
2.1153 | -92.7725
2.1191 | -92.8266
2.1228 | -92.9031
2.1266 | -93.022
2.1303 | -93.2078
2.1341 | -93.4611
2.1378 | -93.7224
2.1416 | -93.8845
2.1453 | -93.8794
2.1491 | -93.7029
2.1528 | -93.4085
2.1566 | -93.1032
2.1603 | -92.8864
2.1641 | -92.8699
2.1678 | -92.8722
2.1716 | -92.9666
2.1753 | -93.1412
2.1791 | -93.3652
2.1828 | -93.5886
2.1866 | -93.7535
2.1903 | -93.8208
2.1941 | -93.7869
2.1978 | -93.6685
2.2016 | -93.5488
2.2053 | -93.4893
2.2091 | -93.495
2.2128 | -93.5377
2.2166 | -93.5771
2.2203 | -93.5732
2.2241 | -93.5127
2.2278 | -93.4094
2.2316 | -93.2928
2.2353 | -93.1918
2.2391 | -93.137
2.2428 | -93.1233
2.2466 | -93.1526
2.2504 | -93.2308
2.2541 | -93.3241
2.2579 | -93.4112
2.2616 | -93.4681
2.2654 | -93.5004
2.2691 | -93.5514
2.2729 | -93.625
2.2766 | -93.6991
2.2804 | -93.7395
2.2841 | -93.738
2.2879 | -93.7315
2.2916 | -93.744
2.2954 | -93.7934
2.2991 | -93.8649
2.3029 | -93.9496
2.3066 | -94.0358
2.3104 | -94.0953
2.3141 | -94.1187
2.3179 | -94.0623
2.3216 | -93.94
2.3254 | -93.8128
2.3291 | -93.7357
2.3329 | -93.733
2.3366 | -93.7827
2.3404 | -93.8616
2.3441 | -93.9635
2.3479 | -94.0878
2.3516 | -94.2314
2.3554 | -94.3583
2.3591 | -94.4367
2.3629 | -94.4433
2.3666 | -94.3786
2.3704 | -94.2616
2.3741 | -94.1084
2.3779 | -93.9137
2.3816 | -93.7225
2.3854 | -93.5834
2.3891 | -93.5243
2.3929 | -93.5275
2.3966 | -93.5627
2.4004 | -93.6394
2.4041 | -93.7541
2.4079 | -93.8978
2.4116 | -94.0552
2.4154 | -94.1978
2.4191 | -94.343
2.4229 | -94.4587
2.4266 | -94.53
2.4304 | -94.5454
2.4341 | -94.5363
2.4379 | -94.5322
2.4416 | -94.5376
2.4454 | -94.5371
2.4491 | -94.4987
2.4529 | -94.4006
2.4566 | -94.2693
2.4604 | -94.1525
2.4641 | -94.0911
2.4679 | -94.0663
2.4716 | -94.0485
2.4754 | -93.9953
2.4791 | -93.9171
2.4829 | -93.8489
2.4866 | -93.8455
2.4904 | -93.9204
2.4941 | -94.0569
2.4979 | -94.2415
2.5016 | -94.4662
2.5054 | -94.7268
2.5091 | -94.9602
2.5129 | -95.0659
2.5166 | -95.0543
2.5204 | -95.0071
2.5241 | -95.0155
2.5279 | -95.1077
2.5316 | -95.2601
2.5354 | -95.4245
2.5391 | -95.5369
2.5429 | -95.5512
2.5466 | -95.4827
2.5504 | -95.3743
2.5541 | -95.26
2.5579 | -95.171
2.5617 | -95.1517
2.5654 | -95.2265
2.5692 | -95.3244
2.5729 | -95.3343
2.5767 | -95.1639
2.5804 | -94.8496
2.5842 | -94.5226
2.5879 | -94.3101
2.5917 | -94.2778
2.5954 | -94.3915
2.5992 | -94.5749
2.6029 | -94.7435
2.6067 | -94.8463
2.6104 | -94.8664
2.6142 | -94.8219
2.6179 | -94.7824
2.6217 | -94.7913
2.6254 | -94.8548
2.6292 | -94.9136
2.6329 | -94.8838
2.6367 | -94.7328
2.6404 | -94.5091
2.6442 | -94.3151
2.6479 | -94.2331
2.6517 | -94.3014
2.6554 | -94.4907
2.6592 | -94.7254
2.6629 | -94.9108
2.6667 | -94.968
2.6704 | -94.8948
2.6742 | -94.7895
2.6779 | -94.7302
2.6817 | -94.7246
2.6854 | -94.7473
2.6892 | -94.7805
2.6929 | -94.8149
2.6967 | -94.8535
2.7004 | -94.8834
2.7042 | -94.9115
2.7079 | -94.9478
2.7117 | -94.9916
2.7154 | -95.0428
2.7192 | -95.0985
2.7229 | -95.1493
2.7267 | -95.1977
2.7304 | -95.2306
2.7342 | -95.2362
2.7379 | -95.2297
2.7417 | -95.2304
2.7454 | -95.2776
2.7492 | -95.366
2.7529 | -95.4602
2.7567 | -95.5207
2.7604 | -95.5122
2.7642 | -95.4297
2.7679 | -95.3049
2.7717 | -95.2021
2.7754 | -95.1518
2.7792 | -95.1672
2.7829 | -95.2464
2.7867 | -95.3448
2.7904 | -95.4394
2.7942 | -95.5074
2.7979 | -95.5431
2.8017 | -95.5648
2.8054 | -95.5896
2.8092 | -95.6222
2.8129 | -95.6401
2.8167 | -95.6132
2.8204 | -95.5437
2.8242 | -95.4679
2.8279 | -95.3992
2.8317 | -95.3329
2.8354 | -95.2855
2.8392 | -95.2629
2.8429 | -95.2968
2.8467 | -95.4009
2.8504 | -95.5451
2.8542 | -95.6838
2.8579 | -95.7755
2.8617 | -95.7892
2.8654 | -95.7378
2.8692 | -95.6491
2.8729 | -95.5494
2.8767 | -95.4921
2.8805 | -95.498
2.8842 | -95.5541
2.888 | -95.6532
2.8917 | -95.7769
2.8955 | -95.9201
2.8992 | -96.0807
2.903 | -96.2344
2.9067 | -96.3399
2.9105 | -96.3543
2.9142 | -96.2475
2.918 | -96.0593
2.9217 | -95.8419
2.9255 | -95.6461
2.9292 | -95.5169
2.933 | -95.5043
2.9367 | -95.5949
2.9405 | -95.7601
2.9442 | -95.9406
2.948 | -96.0639
2.9517 | -96.0928
2.9555 | -96.0229
2.9592 | -95.9015
2.963 | -95.7604
2.9667 | -95.6188
2.9705 | -95.5217
2.9742 | -95.47
2.978 | -95.4629
2.9817 | -95.4908
2.9855 | -95.5762
2.9892 | -95.7291
2.993 | -95.9336
2.9967 | -96.1578
3.0005 | -96.3669
3.0042 | -96.5208
3.008 | -96.598
3.0117 | -96.6418
3.0155 | -96.6818
3.0192 | -96.73
3.023 | -96.827
3.0267 | -96.958
3.0305 | -97.0751
3.0342 | -97.1383
3.038 | -97.12
3.0417 | -97.038
3.0455 | -96.9084
3.0492 | -96.7591
3.053 | -96.6284
3.0567 | -96.5128
3.0605 | -96.4103
3.0642 | -96.3209
3.068 | -96.2421
3.0717 | -96.1629
3.0755 | -96.0878
3.0792 | -96.0273
3.083 | -95.9898
3.0867 | -95.9923
3.0905 | -96.0474
3.0942 | -96.1586
3.098 | -96.3099
3.1017 | -96.4895
3.1055 | -96.6964
3.1092 | -96.903
3.113 | -97.0951
3.1167 | -97.2196
3.1205 | -97.2516
3.1242 | -97.1805
3.128 | -97.0189
3.1317 | -96.8175
3.1355 | -96.6374
3.1392 | -96.4988
3.143 | -96.4154
3.1467 | -96.3643
3.1505 | -96.3168
3.1542 | -96.2765
3.158 | -96.269
3.1617 | -96.2861
3.1655 | -96.3355
3.1692 | -96.3736
3.173 | -96.3775
3.1767 | -96.3513
3.1805 | -96.3126
3.1842 | -96.3022
3.188 | -96.3454
3.1917 | -96.4481
3.1955 | -96.6084
3.1993 | -96.8079
3.203 | -97.0047
3.2068 | -97.1566
3.2105 | -97.2369
3.2143 | -97.2255
3.218 | -97.1717
3.2218 | -97.1145
3.2255 | -97.0944
3.2293 | -97.1249
3.233 | -97.1666
3.2368 | -97.1907
3.2405 | -97.1662
3.2443 | -97.054
3.248 | -96.9699
3.2518 | -96.8924
3.2555 | -96.8342
3.2593 | -96.8003
3.263 | -96.7898
3.2668 | -96.7857
3.2705 | -96.7647
3.2743 | -96.7243
3.278 | -96.6653
3.2818 | -96.5953
3.2855 | -96.5392
3.2893 | -96.5184
3.293 | -96.5359
3.2968 | -96.5923
3.3005 | -96.6668
3.3043 | -96.7465
3.308 | -96.7974
3.3118 | -96.8193
3.3155 | -96.8198
3.3193 | -96.8037
3.323 | -96.7936
3.3268 | -96.8217
3.3305 | -96.8977
3.3343 | -97.0234
3.338 | -97.186
3.3418 | -97.3504
3.3455 | -97.4703
3.3493 | -97.5199
3.353 | -97.4864
3.3568 | -97.3938
3.3605 | -97.2636
3.3643 | -97.1442
3.368 | -97.0657
3.3718 | -97.036
3.3755 | -97.0667
3.3793 | -97.1341
3.383 | -97.2028
3.3868 | -97.2478
3.3905 | -97.2484
3.3943 | -97.2002
3.398 | -97.1407
3.4018 | -97.096
3.4055 | -97.0907
3.4093 | -97.1335
3.413 | -97.2121
3.4168 | -97.3187
3.4205 | -97.4432
3.4243 | -97.565
3.428 | -97.6645
3.4318 | -97.7322
3.4355 | -97.7582
3.4393 | -97.7348
3.443 | -97.6732
3.4468 | -97.5743
3.4505 | -97.4606
3.4543 | -97.3274
3.458 | -97.2057
3.4618 | -97.1032
3.4655 | -97.0121
3.4693 | -96.9424
3.473 | -96.8947
3.4768 | -96.8725
3.4805 | -96.8764
3.4843 | -96.8986
3.488 | -96.9185
3.4918 | -96.943
3.4955 | -96.9784
3.4993 | -97.0145
3.503 | -97.0696
3.5068 | -97.1433
3.5105 | -97.2433
3.5143 | -97.3566
3.5181 | -97.4675
3.5218 | -97.5563
3.5256 | -97.6015
3.5293 | -97.6046
3.5331 | -97.578
3.5368 | -97.5591
3.5406 | -97.5615
3.5443 | -97.599
3.5481 | -97.6543
3.5518 | -97.7083
3.5556 | -97.7613
3.5593 | -97.7999
3.5631 | -97.8145
3.5668 | -97.8028
3.5706 | -97.7569
3.5743 | -97.6767
3.5781 | -97.5671
3.5818 | -97.4405
3.5856 | -97.3192
3.5893 | -97.2317
3.5931 | -97.1993
3.5968 | -97.2359
3.6006 | -97.3331
3.6043 | -97.4743
3.6081 | -97.6184
3.6118 | -97.722
3.6156 | -97.7683
3.6193 | -97.767
3.6231 | -97.7395
3.6268 | -97.7109
3.6306 | -97.706
3.6343 | -97.7412
3.6381 | -97.827
3.6418 | -97.9638
3.6456 | -98.1367
3.6493 | -98.317
3.6531 | -98.4614
3.6568 | -98.5426
3.6606 | -98.5543
3.6643 | -98.5107
3.6681 | -98.4387
3.6718 | -98.3695
3.6756 | -98.3043
3.6793 | -98.2434
3.6831 | -98.1824
3.6868 | -98.1243
3.6906 | -98.08
3.6943 | -98.0418
3.6981 | -98.0203
3.7018 | -98.0119
3.7056 | -98.0053
3.7093 | -98.0039
3.7131 | -97.9979
3.7168 | -97.9969
3.7206 | -97.9979
3.7243 | -98.0082
3.7281 | -98.0172
3.7318 | -98.03
3.7356 | -98.0349
3.7393 | -98.041
3.7431 | -98.0386
3.7468 | -98.0269
3.7506 | -98.0148
3.7543 | -97.9933
3.7581 | -97.975
3.7618 | -97.9961
3.7656 | -98.0631
3.7693 | -98.1704
3.7731 | -98.295
3.7768 | -98.4162
3.7806 | -98.491
3.7843 | -98.4683
3.7881 | -98.4667
3.7918 | -98.4387
3.7956 | -98.3915
3.7993 | -98.3321
3.8031 | -98.2832
3.8068 | -98.2537
3.8106 | -98.2547
3.8143 | -98.2886
3.8181 | -98.347
3.8218 | -98.4094
3.8256 | -98.4643
3.8293 | -98.5067
3.8331 | -98.5343
3.8369 | -98.5503
3.8406 | -98.5616
3.8444 | -98.5727
3.8481 | -98.5865
3.8519 | -98.6055
3.8556 | -98.637
3.8594 | -98.6793
3.8631 | -98.7194
3.8669 | -98.7516
3.8706 | -98.7818
3.8744 | -98.8055
3.8781 | -98.8216
3.8819 | -98.8324
3.8856 | -98.8355
3.8894 | -98.8354
3.8931 | -98.8202
3.8969 | -98.787
3.9006 | -98.7485
3.9044 | -98.7134
3.9081 | -98.6569
3.9119 | -98.5977
3.9156 | -98.5417
3.9194 | -98.4793
3.9231 | -98.4336
3.9269 | -98.4154
3.9306 | -98.4263
3.9344 | -98.4651
3.9381 | -98.5332
3.9419 | -98.6218
3.9456 | -98.716
3.9494 | -98.8097
3.9531 | -98.8913
3.9569 | -98.9518
3.9606 | -98.9901
3.9644 | -99.0045
3.9681 | -98.9989
3.9719 | -98.992
3.9756 | -98.9733
3.9794 | -98.9561
3.9831 | -98.93
3.9869 | -98.883
3.9906 | -98.823
3.9944 | -98.7471
3.9981 | -98.6548
4.0019 | -98.5525
4.0056 | -98.4495
4.0094 | -98.3628
4.0131 | -98.3057
4.0169 | -98.2895
4.0206 | -98.3223
4.0244 | -98.3981
4.0281 | -98.5079
4.0319 | -98.6471
4.0356 | -98.7919
4.0394 | -98.9411
4.0431 | -99.0801
4.0469 | -99.2207
4.0506 | -99.3411
4.0544 | -99.4238
4.0581 | -99.4628
4.0619 | -99.4593
4.0656 | -99.4108
4.0694 | -99.3195
4.0731 | -99.2094
4.0769 | -99.1032
4.0806 | -99.0168
4.0844 | -98.9472
4.0881 | -98.8964
4.0919 | -98.8645
4.0956 | -98.848
4.0994 | -98.849
4.1031 | -98.8532
4.1069 | -98.8556
4.1106 | -98.8437
4.1144 | -98.8265
4.1181 | -98.7995
4.1219 | -98.7754
4.1256 | -98.7601
4.1294 | -98.75
4.1331 | -98.7516
4.1369 | -98.7526
4.1406 | -98.7411
4.1444 | -98.7238
4.1481 | -98.7115
4.1519 | -98.6951
4.1556 | -98.684
4.1594 | -98.6905
4.1631 | -98.724
4.1669 | -98.7843
4.1707 | -98.8815
4.1744 | -99.0135
4.1782 | -99.1572
4.1819 | -99.2937
4.1857 | -99.4052
4.1894 | -99.483
4.1932 | -99.5314
4.1969 | -99.5356
4.2007 | -99.4987
4.2044 | -99.447
4.2082 | -99.3823
4.2119 | -99.3084
4.2157 | -99.2348
4.2194 | -99.1715
4.2232 | -99.1315
4.2269 | -99.1198
4.2307 | -99.1379
4.2344 | -99.181
4.2382 | -99.2435
4.2419 | -99.3182
4.2457 | -99.4005
4.2494 | -99.4736
4.2532 | -99.5325
4.2569 | -99.5673
4.2607 | -99.577
4.2644 | -99.5741
4.2682 | -99.5584
4.2719 | -99.5396
4.2757 | -99.5396
4.2794 | -99.5618
4.2832 | -99.5924
4.2869 | -99.6327
4.2907 | -99.6727
4.2944 | -99.7037
4.2982 | -99.7305
4.3019 | -99.7448
4.3057 | -99.7415
4.3094 | -99.7102
4.3132 | -99.6435
4.3169 | -99.551
4.3207 | -99.4513
4.3244 | -99.344
4.3282 | -99.2387
4.3319 | -99.1507
4.3357 | -99.087
4.3394 | -99.0565
4.3432 | -99.058
4.3469 | -99.0947
4.3507 | -99.1645
4.3544 | -99.2648
4.3582 | -99.3939
4.3619 | -99.5123
4.3657 | -99.601
4.3694 | -99.6659
4.3732 | -99.7018
4.3769 | -99.7142
4.3807 | -99.7088
4.3844 | -99.6954
4.3882 | -99.6748
4.3919 | -99.6625
4.3957 | -99.6632
4.3994 | -99.6635
4.4032 | -99.6564
4.4069 | -99.6472
4.4107 | -99.6283
4.4144 | -99.5984
4.4182 | -99.5637
4.4219 | -99.5311
4.4257 | -99.5072
4.4294 | -99.4886
4.4332 | -99.4954
4.4369 | -99.5086
4.4407 | -99.528
4.4444 | -99.5581
4.4482 | -99.5927
4.4519 | -99.6271
4.4557 | -99.6537
4.4594 | -99.6677
4.4632 | -99.666
4.4669 | -99.6477
4.4707 | -99.6227
4.4744 | -99.5894
4.4782 | -99.5487
4.482 | -99.5101
4.4857 | -99.4771
4.4895 | -99.4536
4.4932 | -99.4353
4.497 | -99.4246
4.5007 | -99.442
4.5045 | -99.4807
4.5082 | -99.5317
4.512 | -99.5935
4.5157 | -99.6556
4.5195 | -99.7236
4.5232 | -99.7848
4.527 | -99.8324
4.5307 | -99.8667
4.5345 | -99.8754
4.5382 | -99.8585
4.542 | -99.8184
4.5457 | -99.7644
4.5495 | -99.7064
4.5532 | -99.6466
4.557 | -99.5911
4.5607 | -99.5475
4.5645 | -99.53
4.5682 | -99.5312
4.572 | -99.5465
4.5757 | -99.5772
4.5795 | -99.6223
4.5832 | -99.6623
4.587 | -99.7061
4.5907 | -99.7484
4.5945 | -99.7969
4.5982 | -99.8495
4.602 | -99.8999
4.6057 | -99.9477
4.6095 | -99.9769
4.6132 | -99.9953
4.617 | -100.0034
4.6207 | -100.0024
4.6245 | -100.0085
4.6282 | -100.026
4.632 | -100.0576
4.6357 | -100.0995
4.6395 | -100.1586
4.6432 | -100.2302
4.647 | -100.3036
4.6507 | -100.366
4.6545 | -100.4173
4.6582 | -100.4365
4.662 | -100.4192
4.6657 | -100.3782
4.6695 | -100.3086
4.6732 | -100.2256
4.677 | -100.136
4.6807 | -100.046
4.6845 | -99.9477
4.6882 | -99.8566
4.692 | -99.7906
4.6957 | -99.7481
4.6995 | -99.7342
4.7032 | -99.7381
4.707 | -99.7666
4.7107 | -99.8118
4.7145 | -99.8575
4.7182 | -99.9047
4.722 | -99.947
4.7257 | -99.984
4.7295 | -100.0216
4.7332 | -100.0497
4.737 | -100.0695
4.7407 | -100.0859
4.7445 | -100.0947
4.7482 | -100.0999
4.752 | -100.1081
4.7557 | -100.1289
4.7595 | -100.1521
4.7632 | -100.1671
4.767 | -100.1698
4.7707 | -100.1527
4.7745 | -100.129
4.7782 | -100.0763
4.782 | -100.0078
4.7857 | -99.9454
4.7895 | -99.8901
4.7932 | -99.8566
4.797 | -99.8486
4.8007 | -99.8613
4.8045 | -99.8992
4.8083 | -99.9584
4.812 | -100.0332
4.8158 | -100.1053
4.8195 | -100.1746
4.8233 | -100.2231
4.827 | -100.2447
4.8308 | -100.2464
4.8345 | -100.2328
4.8383 | -100.209
4.842 | -100.1818
4.8458 | -100.1543
4.8495 | -100.1478
4.8533 | -100.1614
4.857 | -100.1906
4.8608 | -100.2258
4.8645 | -100.2595
4.8683 | -100.2906
4.872 | -100.3106
4.8758 | -100.3256
4.8795 | -100.3327
4.8833 | -100.3255
4.887 | -100.3046
4.8908 | -100.2744
4.8945 | -100.2292
4.8983 | -100.1883
4.902 | -100.1443
4.9058 | -100.1055
4.9095 | -100.0757
4.9133 | -100.0588
4.917 | -100.0621
4.9208 | -100.0815
4.9245 | -100.117
4.9283 | -100.1598
4.932 | -100.2102
4.9358 | -100.2752
4.9395 | -100.3359
4.9433 | -100.3781
4.947 | -100.4189
4.9508 | -100.4362
4.9545 | -100.4406
4.9583 | -100.4253
4.962 | -100.3874
4.9658 | -100.3528
4.9695 | -100.3226
4.9733 | -100.2964
4.977 | -100.2864
4.9808 | -100.3076
4.9845 | -100.348
4.9883 | -100.4028
4.992 | -100.4766
4.9958 | -100.5558
4.9995 | -100.6326
5.0033 | -100.6979
5.007 | -100.7507
5.0108 | -100.7801
5.0145 | -100.7845
5.0183 | -100.7765
5.022 | -100.7624
5.0258 | -100.7372
5.0295 | -100.7312
5.0333 | -100.7326
5.037 | -100.7507
5.0408 | -100.7813
5.0445 | -100.8101
5.0483 | -100.8359
5.052 | -100.8624
5.0558 | -100.8683
5.0595 | -100.8567
5.0633 | -100.84
5.067 | -100.819
5.0708 | -100.7834
5.0745 | -100.741
5.0783 | -100.7027
5.082 | -100.6679
5.0858 | -100.6412
5.0895 | -100.6327
5.0933 | -100.6448
5.097 | -100.6702
5.1008 | -100.7077
5.1045 | -100.761
5.1083 | -100.8146
5.112 | -100.8558
5.1158 | -100.8876
5.1196 | -100.9176
5.1233 | -100.9382
5.1271 | -100.9547
5.1308 | -100.9667
5.1346 | -100.9729
5.1383 | -100.9691
5.1421 | -100.9518
5.1458 | -100.9203
5.1496 | -100.881
5.1533 | -100.835
5.1571 | -100.7934
5.1608 | -100.7581
5.1646 | -100.7293
5.1683 | -100.7125
5.1721 | -100.7094
5.1758 | -100.721
5.1796 | -100.7441
5.1833 | -100.7653
5.1871 | -100.7833
5.1908 | -100.8005
5.1946 | -100.7938
5.1983 | -100.7625
5.2021 | -100.7189
5.2058 | -100.6535
5.2096 | -100.5775
5.2133 | -100.5069
5.2171 | -100.4485
5.2208 | -100.4073
5.2246 | -100.3876
5.2283 | -100.3848
5.2321 | -100.4025
5.2358 | -100.4283
5.2396 | -100.4607
5.2433 | -100.4956
5.2471 | -100.5117
5.2508 | -100.5111
5.2546 | -100.4955
5.2583 | -100.45
5.2621 | -100.3835
5.2658 | -100.3109
5.2696 | -100.2356
5.2733 | -100.1713
5.2771 | -100.124
5.2808 | -100.0962
5.2846 | -100.1037
5.2883 | -100.1383
5.2921 | -100.1911
5.2958 | -100.2599
5.2996 | -100.344
5.3033 | -100.432
5.3071 | -100.5104
5.3108 | -100.5802
5.3146 | -100.6418
5.3183 | -100.6988
5.3221 | -100.7445
5.3258 | -100.7868
5.3296 | -100.8233
5.3333 | -100.8573
5.3371 | -100.8901
5.3408 | -100.9159
5.3446 | -100.9257
5.3483 | -100.925
5.3521 | -100.9153
5.3558 | -100.8923
5.3596 | -100.8634
5.3633 | -100.8303
5.3671 | -100.8005
5.3708 | -100.7736
5.3746 | -100.7447
5.3783 | -100.7196
5.3821 | -100.7087
5.3858 | -100.7138
5.3896 | -100.7178
5.3933 | -100.7283
5.3971 | -100.7414
5.4008 | -100.7548
5.4046 | -100.7675
5.4083 | -100.7316
5.4121 | -100.7094
5.4158 | -100.6746
5.4196 | -100.6314
5.4233 | -100.5832
5.4271 | -100.5302
5.4308 | -100.4831
5.4346 | -100.4553
5.4383 | -100.4428
5.4421 | -100.4477
5.4459 | -100.4734
5.4496 | -100.5168
5.4534 | -100.5796
5.4571 | -100.6577
5.4609 | -100.7385
5.4646 | -100.8277
5.4684 | -100.9262
5.4721 | -101.0359
5.4759 | -101.1504
5.4796 | -101.2581
5.4834 | -101.362
5.4871 | -101.4571
5.4909 | -101.5442
5.4946 | -101.6129
5.4984 | -101.6609
5.5021 | -101.6824
5.5059 | -101.6887
5.5096 | -101.676
5.5134 | -101.6467
5.5171 | -101.602
5.5209 | -101.5545
5.5246 | -101.5118
5.5284 | -101.468
5.5321 | -101.4286
5.5359 | -101.4048
5.5396 | -101.3875
5.5434 | -101.3739
5.5471 | -101.3678
5.5509 | -101.3593
5.5546 | -101.3428
5.5584 | -101.3182
5.5621 | -101.2832
5.5659 | -101.2462
5.5696 | -101.2036
5.5734 | -101.1578
5.5771 | -101.1196
5.5809 | -101.0908
5.5846 | -101.0664
5.5884 | -101.0516
5.5921 | -101.0372
5.5959 | -101.0273
5.5996 | -101.0181
5.6034 | -101.0035
5.6071 | -100.9777
5.6109 | -100.952
5.6146 | -100.9321
5.6184 | -100.9183
5.6221 | -100.9144
5.6259 | -100.9228
5.6296 | -100.9394
5.6334 | -100.9612
5.6371 | -100.9829
5.6409 | -101.0036
5.6446 | -101.0154
5.6484 | -101.0177
5.6521 | -101.0152
5.6559 | -101.0037
5.6596 | -100.9856
5.6634 | -100.9735
5.6671 | -100.973
5.6709 | -100.9865
5.6746 | -101.0062
5.6784 | -101.0361
5.6821 | -101.069
5.6859 | -101.1057
5.6896 | -101.1501
5.6934 | -101.2
5.6971 | -101.2521
5.7009 | -101.3026
5.7046 | -101.3501
5.7084 | -101.3903
5.7121 | -101.4342
5.7159 | -101.4799
5.7196 | -101.5206
5.7234 | -101.5544
5.7271 | -101.5724
5.7309 | -101.5711
5.7346 | -101.5488
5.7384 | -101.5066
5.7421 | -101.4516
5.7459 | -101.3843
5.7496 | -101.2963
5.7534 | -101.2059
5.7572 | -101.1247
5.7609 | -101.0631
5.7647 | -101.0049
5.7684 | -100.9668
5.7722 | -100.9589
5.7759 | -100.9646
5.7797 | -100.9854
5.7834 | -101.019
5.7872 | -101.0671
5.7909 | -101.1246
5.7947 | -101.1827
5.7984 | -101.2389
5.8022 | -101.2906
5.8059 | -101.3432
5.8097 | -101.395
5.8134 | -101.4504
5.8172 | -101.5038
5.8209 | -101.545
5.8247 | -101.5722
5.8284 | -101.5927
5.8322 | -101.6148
5.8359 | -101.6429
5.8397 | -101.6677
5.8434 | -101.6882
5.8472 | -101.7026
5.8509 | -101.7239
5.8547 | -101.7347
5.8584 | -101.7254
5.8622 | -101.7084
5.8659 | -101.675
5.8697 | -101.6391
5.8734 | -101.5981
5.8772 | -101.5578
5.8809 | -101.5382
5.8847 | -101.5377
5.8884 | -101.5597
5.8922 | -101.6122
5.8959 | -101.6916
5.8997 | -101.7923
5.9034 | -101.9069
5.9072 | -102.0232
5.9109 | -102.1295
5.9147 | -102.214
5.9184 | -102.2641
5.9222 | -102.2835
5.9259 | -102.2806
5.9297 | -102.2522
5.9334 | -102.2038
5.9372 | -102.1404
5.9409 | -102.0714
5.9447 | -102.0043
5.9484 | -101.9645
5.9522 | -101.9147
5.9559 | -101.8733
5.9597 | -101.8393
5.9634 | -101.8179
5.9672 | -101.8089
5.9709 | -101.7934
5.9747 | -101.7826
5.9784 | -101.7778
5.9822 | -101.7663
5.9859 | -101.7549
5.9897 | -101.7381
5.9934 | -101.7198
5.9972 | -101.7081
6.0009 | -101.6914
6.0047 | -101.6782
6.0084 | -101.6655
6.0122 | -101.6578
6.0159 | -101.6522
6.0197 | -101.6566
6.0234 | -101.6682
6.0272 | -101.6835
6.0309 | -101.6984
6.0347 | -101.7213
6.0384 | -101.7462
6.0422 | -101.7744
6.0459 | -101.7994
6.0497 | -101.8229
6.0534 | -101.8372
6.0572 | -101.8447
6.0609 | -101.8417
6.0647 | -101.8367
6.0684 | -101.8227
6.0722 | -101.8044
6.0759 | -101.791
6.0797 | -101.778
6.0835 | -101.7699
6.0872 | -101.7744
6.091 | -101.784
6.0947 | -101.8008
6.0985 | -101.8173
6.1022 | -101.8381
6.106 | -101.8542
6.1097 | -101.8629
6.1135 | -101.8626
6.1172 | -101.8619
6.121 | -101.8627
6.1247 | -101.8707
6.1285 | -101.8715
6.1322 | -101.8716
6.136 | -101.8723
6.1397 | -101.8742
6.1435 | -101.8737
6.1472 | -101.8747
6.151 | -101.8796
6.1547 | -101.8899
6.1585 | -101.8941
6.1622 | -101.8971
6.166 | -101.9021
6.1697 | -101.9085
6.1735 | -101.9223
6.1772 | -101.9421
6.181 | -101.9557
6.1847 | -101.9644
6.1885 | -101.9758
6.1922 | -101.9855
6.196 | -101.9871
6.1997 | -101.9829
6.2035 | -101.97
6.2072 | -101.9487
6.211 | -101.9153
6.2147 | -101.8756
6.2185 | -101.8319
6.2222 | -101.7976
6.226 | -101.7632
6.2297 | -101.7346
6.2335 | -101.7183
6.2372 | -101.713
6.241 | -101.721
6.2447 | -101.7445
6.2485 | -101.7825
6.2522 | -101.8417
6.256 | -101.9074
6.2597 | -101.9667
6.2635 | -102.0267
6.2672 | -102.0852
6.271 | -102.1296
6.2747 | -102.1619
6.2785 | -102.1829
6.2822 | -102.1866
6.286 | -102.171
6.2897 | -102.1465
6.2935 | -102.1152
6.2972 | -102.0828
6.301 | -102.0437
6.3047 | -101.9967
6.3085 | -101.945
6.3122 | -101.8884
6.316 | -101.8272
6.3197 | -101.7718
6.3235 | -101.7174
6.3272 | -101.6673
6.331 | -101.6324
6.3347 | -101.6067
6.3385 | -101.5951
6.3422 | -101.6047
6.346 | -101.6277
6.3497 | -101.6711
6.3535 | -101.7325
6.3572 | -101.7995
6.361 | -101.8703
6.3647 | -101.9482
6.3685 | -102.028
6.3722 | -102.1126
6.376 | -102.1898
6.3797 | -102.2618
6.3835 | -102.3277
6.3872 | -102.3807
6.391 | -102.4179
6.3947 | -102.4515
6.3985 | -102.4783
6.4023 | -102.4893
6.406 | -102.4981
6.4098 | -102.5046
6.4135 | -102.5035
6.4173 | -102.4988
6.421 | -102.4922
6.4248 | -102.4811
6.4285 | -102.4688
6.4323 | -102.4431
6.436 | -102.4086
6.4398 | -102.3697
6.4435 | -102.3243
6.4473 | -102.2824
6.451 | -102.2417
6.4548 | -102.2029
6.4585 | -102.1686
6.4623 | -102.1401
6.466 | -102.1187
6.4698 | -102.1062
6.4735 | -102.0974
6.4773 | -102.0989
6.481 | -102.1098
6.4848 | -102.1299
6.4885 | -102.1503
6.4923 | -102.1676
6.496 | -102.1802
6.4998 | -102.1907
6.5035 | -102.1952
6.5073 | -102.2003
6.511 | -102.2111
6.5148 | -102.2289
6.5185 | -102.2512
6.5223 | -102.2735
6.526 | -102.2993
6.5298 | -102.3275
6.5335 | -102.3506
6.5373 | -102.3698
6.541 | -102.3734
6.5448 | -102.3685
6.5485 | -102.3594
6.5523 | -102.3441
6.556 | -102.3248
6.5598 | -102.309
6.5635 | -102.2893
6.5673 | -102.2724
6.571 | -102.2594
6.5748 | -102.2459
6.5785 | -102.2374
6.5823 | -102.2353
6.586 | -102.2343
6.5898 | -102.244
6.5935 | -102.2589
6.5973 | -102.2802
6.601 | -102.3053
6.6048 | -102.3335
6.6085 | -102.357
6.6123 | -102.3814
6.616 | -102.3997
6.6198 | -102.4205
6.6235 | -102.4456
6.6273 | -102.467
6.631 | -102.4823
6.6348 | -102.4917
6.6385 | -102.4902
6.6423 | -102.4791
6.646 | -102.4603
6.6498 | -102.4349
6.6535 | -102.4017
6.6573 | -102.3639
6.661 | -102.3196
6.6648 | -102.2789
6.6685 | -102.2388
6.6723 | -102.2042
6.676 | -102.1675
6.6798 | -102.1407
6.6835 | -102.1259
6.6873 | -102.1264
6.691 | -102.1315
6.6948 | -102.1481
6.6985 | -102.182
6.7023 | -102.2271
6.706 | -102.2761
6.7098 | -102.3341
6.7135 | -102.3926
6.7173 | -102.4541
6.721 | -102.5128
6.7248 | -102.5591
6.7286 | -102.5968
6.7323 | -102.6283
6.7361 | -102.6429
6.7398 | -102.6494
6.7436 | -102.6534
6.7473 | -102.6406
6.7511 | -102.6153
6.7548 | -102.591
6.7586 | -102.5667
6.7623 | -102.5445
6.7661 | -102.5203
6.7698 | -102.499
6.7736 | -102.4783
6.7773 | -102.4639
6.7811 | -102.4419
6.7848 | -102.4197
6.7886 | -102.3973
6.7923 | -102.38
6.7961 | -102.3657
6.7998 | -102.3482
6.8036 | -102.3284
6.8073 | -102.3057
6.8111 | -102.2906
6.8148 | -102.2785
6.8186 | -102.2567
6.8223 | -102.2398
6.8261 | -102.2321
6.8298 | -102.2311
6.8336 | -102.231
6.8373 | -102.2351
6.8411 | -102.2392
6.8448 | -102.2505
6.8486 | -102.269
6.8523 | -102.2905
6.8561 | -102.3106
6.8598 | -102.3339
6.8636 | -102.3649
6.8673 | -102.3997
6.8711 | -102.4368
6.8748 | -102.4791
6.8786 | -102.5265
6.8823 | -102.5806
6.8861 | -102.6342
6.8898 | -102.6936
6.8936 | -102.7526
6.8973 | -102.808
6.9011 | -102.8523
6.9048 | -102.8993
6.9086 | -102.9411
6.9123 | -102.9655
6.9161 | -102.9803
6.9198 | -102.9887
6.9236 | -102.978
6.9273 | -102.9608
6.9311 | -102.929
6.9348 | -102.8926
6.9386 | -102.8537
6.9423 | -102.8158
6.9461 | -102.7733
6.9498 | -102.7329
6.9536 | -102.6954
6.9573 | -102.6576
6.9611 | -102.6215
6.9648 | -102.5888
6.9686 | -102.5668
6.9723 | -102.5516
6.9761 | -102.5459
6.9798 | -102.5497
6.9836 | -102.5575
6.9873 | -102.5703
6.9911 | -102.5901
6.9948 | -102.6189
6.9986 | -102.6603
7.0023 | -102.7179
7.0061 | -102.7813
7.0098 | -102.8528
7.0136 | -102.9367
7.0173 | -103.0306
7.0211 | -103.1355
7.0248 | -103.2397
7.0286 | -103.3331
7.0323 | -103.3766
7.0361 | -103.4309
7.0399 | -103.4702
7.0436 | -103.4869
7.0474 | -103.4842
7.0511 | -103.4611
7.0549 | -103.4159
7.0586 | -103.3577
7.0624 | -103.2934
7.0661 | -103.2208
7.0699 | -103.1345
7.0736 | -103.0466
7.0774 | -102.9615
7.0811 | -102.8768
7.0849 | -102.799
7.0886 | -102.7353
7.0924 | -102.6799
7.0961 | -102.6341
7.0999 | -102.6005
7.1036 | -102.5761
7.1074 | -102.5665
7.1111 | -102.5667
7.1149 | -102.5789
7.1186 | -102.5971
7.1224 | -102.6199
7.1261 | -102.6389
7.1299 | -102.6602
7.1336 | -102.68
7.1374 | -102.6962
7.1411 | -102.7151
7.1449 | -102.7356
7.1486 | -102.7515
7.1524 | -102.7675
7.1561 | -102.7795
7.1599 | -102.7946
7.1636 | -102.8171
7.1674 | -102.8422
7.1711 | -102.863
7.1749 | -102.8788
7.1786 | -102.8977
7.1824 | -102.9116
7.1861 | -102.9205
7.1899 | -102.9284
7.1936 | -102.9367
7.1974 | -102.9481
7.2011 | -102.9653
7.2049 | -102.9813
7.2086 | -102.9951
7.2124 | -103.0143
7.2161 | -103.0265
7.2199 | -103.0415
7.2236 | -103.0575
7.2274 | -103.0633
7.2311 | -103.064
7.2349 | -103.0678
7.2386 | -103.0638
7.2424 | -103.0581
7.2461 | -103.0495
7.2499 | -103.0311
7.2536 | -103.0004
7.2574 | -102.9663
7.2611 | -102.9306
7.2649 | -102.8974
7.2686 | -102.8651
7.2724 | -102.8429
7.2761 | -102.8209
7.2799 | -102.8068
7.2836 | -102.8093
7.2874 | -102.8218
7.2911 | -102.8392
7.2949 | -102.8606
7.2986 | -102.8865
7.3024 | -102.9118
7.3061 | -102.9326
7.3099 | -102.9488
7.3136 | -102.9643
7.3174 | -102.9758
7.3211 | -102.9813
7.3249 | -102.9887
7.3286 | -102.99
7.3324 | -102.9926
7.3361 | -103.0017
7.3399 | -103.0126
7.3436 | -103.018
7.3474 | -103.0267
7.3511 | -103.0375
7.3549 | -103.049
7.3586 | -103.0648
7.3624 | -103.0799
7.3662 | -103.0952
7.3699 | -103.1133
7.3737 | -103.1324
7.3774 | -103.1504
7.3812 | -103.1562
7.3849 | -103.1657
7.3887 | -103.1749
7.3924 | -103.1768
7.3962 | -103.1737
7.3999 | -103.1713
7.4037 | -103.168
7.4074 | -103.1628
7.4112 | -103.1599
7.4149 | -103.1599
7.4187 | -103.1554
7.4224 | -103.151
7.4262 | -103.1534
7.4299 | -103.1511
7.4337 | -103.1534
7.4374 | -103.1633
7.4412 | -103.1725
7.4449 | -103.1863
7.4487 | -103.2043
7.4524 | -103.227
7.4562 | -103.2463
7.4599 | -103.2616
7.4637 | -103.2674
7.4674 | -103.2748
7.4712 | -103.2779
7.4749 | -103.2746
7.4787 | -103.2665
7.4824 | -103.2487
7.4862 | -103.2296
7.4899 | -103.2126
7.4937 | -103.1915
7.4974 | -103.1692
7.5012 | -103.1462
7.5049 | -103.1277
7.5087 | -103.1155
7.5124 | -103.1117
7.5162 | -103.1135
7.5199 | -103.121
7.5237 | -103.1357
7.5274 | -103.1636
7.5312 | -103.1932
7.5349 | -103.2268
7.5387 | -103.2635
7.5424 | -103.2931
7.5462 | -103.3214
7.5499 | -103.3398
7.5537 | -103.3516
7.5574 | -103.3556
7.5612 | -103.3508
7.5649 | -103.3452
7.5687 | -103.336
7.5724 | -103.3174
7.5762 | -103.3107
7.5799 | -103.2946
7.5837 | -103.2731
7.5874 | -103.2564
7.5912 | -103.2364
7.5949 | -103.2244
7.5987 | -103.212
7.6024 | -103.2018
7.6062 | -103.1918
7.6099 | -103.181
7.6137 | -103.1756
7.6174 | -103.1751
7.6212 | -103.1739
7.6249 | -103.1745
7.6287 | -103.1852
7.6324 | -103.2022
7.6362 | -103.2223
7.6399 | -103.2436
7.6437 | -103.2687
7.6474 | -103.2911
7.6512 | -103.3189
7.6549 | -103.3395
7.6587 | -103.3504
7.6624 | -103.3678
7.6662 | -103.3803
7.6699 | -103.3841
7.6737 | -103.3838
7.6775 | -103.3756
7.6812 | -103.3644
7.685 | -103.3525
7.6887 | -103.332
7.6925 | -103.3077
7.6962 | -103.2813
7.7 | -103.2563
7.7037 | -103.2256
7.7075 | -103.194
7.7112 | -103.161
7.715 | -103.1293
7.7187 | -103.0933
7.7225 | -103.0682
7.7262 | -103.0457
7.73 | -103.0204
7.7337 | -102.9978
7.7375 | -102.987
7.7412 | -102.9805
7.745 | -102.9773
7.7487 | -102.9806
7.7525 | -102.9909
7.7562 | -103.0014
7.76 | -103.0172
7.7637 | -103.0394
7.7675 | -103.0673
7.7712 | -103.1073
7.775 | -103.155
7.7787 | -103.203
7.7825 | -103.2553
7.7862 | -103.3105
7.79 | -103.364
7.7937 | -103.4188
7.7975 | -103.4692
7.8012 | -103.5104
7.805 | -103.5518
7.8087 | -103.5846
7.8125 | -103.6047
7.8162 | -103.608
7.82 | -103.5914
7.8237 | -103.5686
7.8275 | -103.5337
7.8312 | -103.4916
7.835 | -103.4499
7.8387 | -103.3939
7.8425 | -103.3362
7.8462 | -103.2815
7.85 | -103.2335
7.8537 | -103.194
7.8575 | -103.1645
7.8612 | -103.1396
7.865 | -103.1286
7.8687 | -103.1289
7.8725 | -103.1453
7.8762 | -103.1684
7.88 | -103.1942
7.8837 | -103.2182
7.8875 | -103.2485
7.8912 | -103.2776
7.895 | -103.3029
7.8987 | -103.3228
7.9025 | -103.3409
7.9062 | -103.3547
7.91 | -103.358
7.9137 | -103.3528
7.9175 | -103.349
7.9212 | -103.3381
7.925 | -103.3281
7.9287 | -103.3175
7.9325 | -103.3119
7.9362 | -103.3177
7.94 | -103.3263
7.9437 | -103.3381
7.9475 | -103.3563
7.9512 | -103.3796
7.955 | -103.4103
7.9587 | -103.4477
7.9625 | -103.4904
7.9662 | -103.5277
7.97 | -103.5567
7.9737 | -103.5902
7.9775 | -103.6268
7.9812 | -103.6596
7.985 | -103.6858
7.9887 | -103.705
7.9925 | -103.7161
7.9962 | -103.7295
8 | -103.7438
8.0038 | -103.7506
8.0075 | -103.7529
8.0113 | -103.7502
8.015 | -103.7452
8.0188 | -103.7336
8.0225 | -103.7166
8.0263 | -103.7036
8.03 | -103.6912
8.0338 | -103.6802
8.0375 | -103.6771
8.0413 | -103.6775
8.045 | -103.6855
8.0488 | -103.6947
8.0525 | -103.7079
8.0563 | -103.724
8.06 | -103.7481
8.0638 | -103.7681
8.0675 | -103.7805
8.0713 | -103.7958
8.075 | -103.8129
8.0788 | -103.8291
8.0825 | -103.8426
8.0863 | -103.8586
8.09 | -103.8653
8.0938 | -103.8646
8.0975 | -103.8603
8.1013 | -103.86
8.105 | -103.8565
8.1088 | -103.8435
8.1125 | -103.8169
8.1163 | -103.7854
8.12 | -103.7612
8.1238 | -103.7345
8.1275 | -103.7098
8.1313 | -103.6842
8.135 | -103.658
8.1388 | -103.6337
8.1425 | -103.6117
8.1463 | -103.5867
8.15 | -103.5638
8.1538 | -103.5413
8.1575 | -103.5169
8.1613 | -103.4929
8.165 | -103.4719
8.1688 | -103.4549
8.1725 | -103.4357
8.1763 | -103.4201
8.18 | -103.4024
8.1838 | -103.3917
8.1875 | -103.3858
8.1913 | -103.3789
8.195 | -103.3704
8.1988 | -103.3664
8.2025 | -103.369
8.2063 | -103.3702
8.21 | -103.3795
8.2138 | -103.3955
8.2175 | -103.4142
8.2213 | -103.4339
8.225 | -103.4612
8.2288 | -103.494
8.2325 | -103.5303
8.2363 | -103.5623
8.24 | -103.5936
8.2438 | -103.6258
8.2475 | -103.6597
8.2513 | -103.6927
8.255 | -103.7263
8.2588 | -103.7522
8.2625 | -103.7741
8.2663 | -103.7869
8.27 | -103.788
8.2738 | -103.7839
8.2775 | -103.7673
8.2813 | -103.7467
8.285 | -103.7263
8.2888 | -103.7005
8.2925 | -103.671
8.2963 | -103.6403
8.3 | -103.6109
8.3038 | -103.5849
8.3075 | -103.563
8.3113 | -103.5439
8.315 | -103.5255
8.3188 | -103.5101
8.3225 | -103.4957
8.3263 | -103.4866
8.3301 | -103.4818
8.3338 | -103.4768
8.3376 | -103.4715
8.3413 | -103.4709
8.3451 | -103.4778
8.3488 | -103.4873
8.3526 | -103.5024
8.3563 | -103.5228
8.3601 | -103.5474
8.3638 | -103.5736
8.3676 | -103.6034
8.3713 | -103.6374
8.3751 | -103.6738
8.3788 | -103.7135
8.3826 | -103.7587
8.3863 | -103.8038
8.3901 | -103.8501
8.3938 | -103.9006
8.3976 | -103.946
8.4013 | -103.9876
8.4051 | -104.0281
8.4088 | -104.0569
8.4126 | -104.0815
8.4163 | -104.1019
8.4201 | -104.1078
8.4238 | -104.1098
8.4276 | -104.113
8.4313 | -104.104
8.4351 | -104.0919
8.4388 | -104.0755
8.4426 | -104.0489
8.4463 | -104.0245
8.4501 | -103.9966
8.4538 | -103.9672
8.4576 | -103.9359
8.4613 | -103.9094
8.4651 | -103.8841
8.4688 | -103.8598
8.4726 | -103.8419
8.4763 | -103.8387
8.4801 | -103.8391
8.4838 | -103.8502
8.4876 | -103.8709
8.4913 | -103.898
8.4951 | -103.9304
8.4988 | -103.9668
8.5026 | -104.0038
8.5063 | -104.0452
8.5101 | -104.0864
8.5138 | -104.1205
8.5176 | -104.1503
8.5213 | -104.1761
8.5251 | -104.1961
8.5288 | -104.2014
8.5326 | -104.2005
8.5363 | -104.1874
8.5401 | -104.1683
8.5438 | -104.1484
8.5476 | -104.1125
8.5513 | -104.0722
8.5551 | -104.0343
8.5588 | -103.9957
8.5626 | -103.9621
8.5663 | -103.9371
8.5701 | -103.9095
8.5738 | -103.8895
8.5776 | -103.8769
8.5813 | -103.8672
8.5851 | -103.863
8.5888 | -103.864
8.5926 | -103.8706
8.5963 | -103.8809
8.6001 | -103.8873
8.6038 | -103.8994
8.6076 | -103.9133
8.6113 | -103.9277
8.6151 | -103.9429
8.6188 | -103.9652
8.6226 | -103.9911
8.6263 | -104.0125
8.6301 | -104.0264
8.6338 | -104.0393
8.6376 | -104.057
8.6413 | -104.0698
8.6451 | -104.0714
8.6489 | -104.0708
8.6526 | -104.0696
8.6564 | -104.0575
8.6601 | -104.0487
8.6639 | -104.0359
8.6676 | -104.0164
8.6714 | -103.9941
8.6751 | -103.9703
8.6789 | -103.9436
8.6826 | -103.9205
8.6864 | -103.8935
8.6901 | -103.8627
8.6939 | -103.8337
8.6976 | -103.8063
8.7014 | -103.7803
8.7051 | -103.7577
8.7089 | -103.7368
8.7126 | -103.7159
8.7164 | -103.6998
8.7201 | -103.6934
8.7239 | -103.6976
8.7276 | -103.7085
8.7314 | -103.7241
8.7351 | -103.7483
8.7389 | -103.78
8.7426 | -103.8177
8.7464 | -103.8552
8.7501 | -103.899
8.7539 | -103.957
8.7576 | -104.0148
8.7614 | -104.0655
8.7651 | -104.1152
8.7689 | -104.1588
8.7726 | -104.1997
8.7764 | -104.2291
8.7801 | -104.2444
8.7839 | -104.2468
8.7876 | -104.242
8.7914 | -104.2296
8.7951 | -104.2058
8.7989 | -104.1741
8.8026 | -104.1356
8.8064 | -104.0919
8.8101 | -104.042
8.8139 | -103.9917
8.8176 | -103.9495
8.8214 | -103.911
8.8251 | -103.8752
8.8289 | -103.8415
8.8326 | -103.81
8.8364 | -103.7878
8.8401 | -103.7714
8.8439 | -103.7567
8.8476 | -103.7446
8.8514 | -103.734
8.8551 | -103.7252
8.8589 | -103.7221
8.8626 | -103.7267
8.8664 | -103.7326
8.8701 | -103.7453
8.8739 | -103.7625
8.8776 | -103.7841
8.8814 | -103.805
8.8851 | -103.8219
8.8889 | -103.843
8.8926 | -103.8708
8.8964 | -103.8935
8.9001 | -103.9176
8.9039 | -103.9436
8.9076 | -103.9692
8.9114 | -103.9986
8.9151 | -104.0282
8.9189 | -104.0547
8.9226 | -104.0815
8.9264 | -104.1106
8.9301 | -104.1398
8.9339 | -104.1726
8.9376 | -104.2059
8.9414 | -104.2391
8.9451 | -104.266
8.9489 | -104.2908
8.9526 | -104.3165
8.9564 | -104.3351
8.9602 | -104.3545
8.9639 | -104.3691
8.9677 | -104.3735
8.9714 | -104.3791
8.9752 | -104.3823
8.9789 | -104.3826
8.9827 | -104.3768
8.9864 | -104.3655
8.9902 | -104.3512
8.9939 | -104.3357
8.9977 | -104.323
9.0014 | -104.3086
9.0052 | -104.2982
9.0089 | -104.2925
9.0127 | -104.2853
9.0164 | -104.283
9.0202 | -104.2845
9.0239 | -104.2875
9.0277 | -104.2962
9.0314 | -104.3071
9.0352 | -104.3172
9.0389 | -104.3265
9.0427 | -104.3359
9.0464 | -104.3495
9.0502 | -104.3669
9.0539 | -104.3822
9.0577 | -104.3996
9.0614 | -104.4136
9.0652 | -104.4184
9.0689 | -104.4226
9.0727 | -104.4333
9.0764 | -104.4432
9.0802 | -104.4487
9.0839 | -104.4566
9.0877 | -104.4672
9.0914 | -104.4746
9.0952 | -104.4809
9.0989 | -104.4894
9.1027 | -104.4907
9.1064 | -104.4908
9.1102 | -104.4913
9.1139 | -104.4937
9.1177 | -104.4984
9.1214 | -104.5021
9.1252 | -104.4986
9.1289 | -104.4926
9.1327 | -104.4882
9.1364 | -104.4817
9.1402 | -104.4725
9.1439 | -104.4733
9.1477 | -104.4725
9.1514 | -104.4662
9.1552 | -104.4568
9.1589 | -104.451
9.1627 | -104.4453
9.1664 | -104.442
9.1702 | -104.4388
9.1739 | -104.4368
9.1777 | -104.4326
9.1814 | -104.4259
9.1852 | -104.4214
9.1889 | -104.4118
9.1927 | -104.4045
9.1964 | -104.3881
9.2002 | -104.3796
9.2039 | -104.3687
9.2077 | -104.3502
9.2114 | -104.3369
9.2152 | -104.3243
9.2189 | -104.3033
9.2227 | -104.2814
9.2264 | -104.2611
9.2302 | -104.2434
9.2339 | -104.2233
9.2377 | -104.2073
9.2414 | -104.1904
9.2452 | -104.1748
9.2489 | -104.1627
9.2527 | -104.1518
9.2564 | -104.1431
9.2602 | -104.1369
9.2639 | -104.1302
9.2677 | -104.1268
9.2714 | -104.1274
9.2752 | -104.1313
9.279 | -104.1316
9.2827 | -104.1334
9.2865 | -104.1417
9.2902 | -104.1536
9.294 | -104.1645
9.2977 | -104.1776
9.3015 | -104.1914
9.3052 | -104.2116
9.309 | -104.2318
9.3127 | -104.2549
9.3165 | -104.2873
9.3202 | -104.3248
9.324 | -104.3672
9.3277 | -104.4108
9.3315 | -104.4582
9.3352 | -104.509
9.339 | -104.561
9.3427 | -104.6117
9.3465 | -104.6624
9.3502 | -104.7076
9.354 | -104.7541
9.3577 | -104.7988
9.3615 | -104.8326
9.3652 | -104.8545
9.369 | -104.8606
9.3727 | -104.8613
9.3765 | -104.8564
9.3802 | -104.8416
9.384 | -104.8162
9.3877 | -104.7815
9.3915 | -104.7425
9.3952 | -104.7012
9.399 | -104.6587
9.4027 | -104.6157
9.4065 | -104.5679
9.4102 | -104.5237
9.414 | -104.4837
9.4177 | -104.4504
9.4215 | -104.424
9.4252 | -104.3997
9.429 | -104.3822
9.4327 | -104.3715
9.4365 | -104.3698
9.4402 | -104.3704
9.444 | -104.3679
9.4477 | -104.369
9.4515 | -104.3732
9.4552 | -104.374
9.459 | -104.3754
9.4627 | -104.376
9.4665 | -104.3735
9.4702 | -104.3685
9.474 | -104.3573
9.4777 | -104.3435
9.4815 | -104.3293
9.4852 | -104.307
9.489 | -104.2831
9.4927 | -104.2576
9.4965 | -104.2347
9.5002 | -104.2147
9.504 | -104.1957
9.5077 | -104.1774
9.5115 | -104.1594
9.5152 | -104.1434
9.519 | -104.1322
9.5227 | -104.1254
9.5265 | -104.1269
9.5302 | -104.1342
9.534 | -104.1505
9.5377 | -104.1741
9.5415 | -104.2031
9.5452 | -104.2366
9.549 | -104.2738
9.5527 | -104.3127
9.5565 | -104.3527
9.5602 | -104.3995
9.564 | -104.4453
9.5677 | -104.485
9.5715 | -104.5222
9.5752 | -104.5536
9.579 | -104.586
9.5827 | -104.6176
9.5865 | -104.6382
9.5902 | -104.6462
9.594 | -104.6492
9.5977 | -104.6478
9.6015 | -104.6432
9.6053 | -104.6357
9.609 | -104.6238
9.6128 | -104.6087
9.6165 | -104.5893
9.6203 | -104.5667
9.624 | -104.5495
9.6278 | -104.5313
9.6315 | -104.5088
9.6353 | -104.4852
9.639 | -104.4654
9.6428 | -104.4507
9.6465 | -104.4407
9.6503 | -104.4287
9.654 | -104.4182
9.6578 | -104.412
9.6615 | -104.4083
9.6653 | -104.4043
9.669 | -104.399
9.6728 | -104.3949
9.6765 | -104.396
9.6803 | -104.3998
9.684 | -104.403
9.6878 | -104.4023
9.6915 | -104.4012
9.6953 | -104.398
9.699 | -104.3876
9.7028 | -104.3777
9.7065 | -104.3697
9.7103 | -104.3601
9.714 | -104.3452
9.7178 | -104.3271
9.7215 | -104.3098
9.7253 | -104.2967
9.729 | -104.2884
9.7328 | -104.2782
9.7365 | -104.274
9.7403 | -104.2642
9.744 | -104.2562
9.7478 | -104.2503
9.7515 | -104.246
9.7553 | -104.2448
9.759 | -104.243
9.7628 | -104.242
9.7665 | -104.2408
9.7703 | -104.2361
9.774 | -104.2318
9.7778 | -104.2304
9.7815 | -104.2283
9.7853 | -104.223
9.789 | -104.2176
9.7928 | -104.2117
9.7965 | -104.2074
9.8003 | -104.2057
9.804 | -104.2056
9.8078 | -104.2075
9.8115 | -104.2121
9.8153 | -104.2126
9.819 | -104.2143
9.8228 | -104.2199
9.8265 | -104.2319
9.8303 | -104.2465
9.834 | -104.2643
9.8378 | -104.2829
9.8415 | -104.3035
9.8453 | -104.3296
9.849 | -104.3542
9.8528 | -104.3757
9.8565 | -104.4042
9.8603 | -104.4311
9.864 | -104.4555
9.8678 | -104.4799
9.8715 | -104.5031
9.8753 | -104.5246
9.879 | -104.5471
9.8828 | -104.572
9.8865 | -104.5951
9.8903 | -104.6149
9.894 | -104.6336
9.8978 | -104.6511
9.9015 | -104.6677
9.9053 | -104.6841
9.909 | -104.6952
9.9128 | -104.7051
9.9165 | -104.7165
9.9203 | -104.7254
9.9241 | -104.7299
9.9278 | -104.7314
9.9316 | -104.7312
9.9353 | -104.7307
9.9391 | -104.7309
9.9428 | -104.7255
9.9466 | -104.7203
9.9503 | -104.7171
9.9541 | -104.7099
9.9578 | -104.7034
9.9616 | -104.7005
9.9653 | -104.6996
9.9691 | -104.6982
9.9728 | -104.6936
9.9766 | -104.6879
9.9803 | -104.6813
9.9841 | -104.6753
9.9878 | -104.6716
9.9916 | -104.6647
9.9953 | -104.6613
9.9991 | -104.6596
10.0028 | -104.6565
10.0066 | -104.6591
10.0103 | -104.6581
10.0141 | -104.6543
10.0178 | -104.6519
10.0216 | -104.6525
10.0253 | -104.6507
10.0291 | -104.6494
10.0328 | -104.6458
10.0366 | -104.643
10.0403 | -104.645
10.0441 | -104.6475
10.0478 | -104.6495
10.0516 | -104.6519
10.0553 | -104.6552
10.0591 | -104.6601
10.0628 | -104.6679
10.0666 | -104.6731
10.0703 | -104.677
10.0741 | -104.6826
10.0778 | -104.6894
10.0816 | -104.6956
10.0853 | -104.7004
10.0891 | -104.7104
10.0928 | -104.7198
10.0966 | -104.7246
10.1003 | -104.7333
10.1041 | -104.7391
10.1078 | -104.7437
10.1116 | -104.7459
10.1153 | -104.7473
10.1191 | -104.7485
10.1228 | -104.7512
10.1266 | -104.7495
10.1303 | -104.7471
10.1341 | -104.7434
10.1378 | -104.7392
10.1416 | -104.7357
10.1453 | -104.7379
10.1491 | -104.7422
10.1528 | -104.7471
10.1566 | -104.7579
10.1603 | -104.7752
10.1641 | -104.7937
10.1678 | -104.8128
10.1716 | -104.8365
10.1753 | -104.8617
10.1791 | -104.8915
10.1828 | -104.9183
10.1866 | -104.9396
10.1903 | -104.9578
10.1941 | -104.9781
10.1978 | -104.9937
10.2016 | -105.0017
10.2053 | -105.0063
10.2091 | -105.0106
10.2128 | -105.0156
10.2166 | -105.0187
10.2203 | -105.0162
10.2241 | -105.0099
10.2278 | -105.0034
10.2316 | -104.9987
10.2353 | -104.9921
10.2391 | -104.9864
10.2429 | -104.9868
10.2466 | -104.9867
10.2504 | -104.9916
10.2541 | -104.9962
10.2579 | -105.0042
10.2616 | -105.0182
10.2654 | -105.034
10.2691 | -105.0499
10.2729 | -105.0699
10.2766 | -105.0909
10.2804 | -105.1076
10.2841 | -105.1226
10.2879 | -105.1385
10.2916 | -105.1527
10.2954 | -105.1612
10.2991 | -105.1705
10.3029 | -105.175
10.3066 | -105.1748
10.3104 | -105.17
10.3141 | -105.1651
10.3179 | -105.1574
10.3216 | -105.1425
10.3254 | -105.125
10.3291 | -105.1084
10.3329 | -105.092
10.3366 | -105.0754
10.3404 | -105.0605
10.3441 | -105.0423
10.3479 | -105.026
10.3516 | -105.0156
10.3554 | -105.0012
10.3591 | -104.9881
10.3629 | -104.9801
10.3666 | -104.9718
10.3704 | -104.9694
10.3741 | -104.9735
10.3779 | -104.9799
10.3816 | -104.9828
10.3854 | -104.9853
10.3891 | -104.9935
10.3929 | -105.0032
10.3966 | -105.0139
10.4004 | -105.0255
10.4041 | -105.0374
10.4079 | -105.05
10.4116 | -105.0613
10.4154 | -105.0726
10.4191 | -105.0779
10.4229 | -105.0821
10.4266 | -105.0828
10.4304 | -105.0844
10.4341 | -105.0833
10.4379 | -105.0793
10.4416 | -105.0727
10.4454 | -105.0644
10.4491 | -105.0565
10.4529 | -105.0445
10.4566 | -105.0325
10.4604 | -105.0225
10.4641 | -105.0106
10.4679 | -104.9962
10.4716 | -104.9823
10.4754 | -104.9716
10.4791 | -104.9598
10.4829 | -104.9474
10.4866 | -104.9363
10.4904 | -104.9238
10.4941 | -104.9106
10.4979 | -104.9003
10.5016 | -104.8912
10.5054 | -104.8813
10.5091 | -104.876
10.5129 | -104.8684
10.5166 | -104.8621
10.5204 | -104.8601
10.5241 | -104.8612
10.5279 | -104.8637
10.5316 | -104.8688
10.5354 | -104.8808
10.5391 | -104.8952
10.5429 | -104.9077
10.5466 | -104.9225
10.5504 | -104.9431
10.5541 | -104.9633
10.5579 | -104.9837
10.5617 | -105.0023
10.5654 | -105.0207
10.5692 | -105.0372
10.5729 | -105.0551
10.5767 | -105.0698
10.5804 | -105.0824
10.5842 | -105.0919
10.5879 | -105.1005
10.5917 | -105.1111
10.5954 | -105.1232
10.5992 | -105.1306
10.6029 | -105.1362
10.6067 | -105.144
10.6104 | -105.1531
10.6142 | -105.1616
10.6179 | -105.1691
10.6217 | -105.1729
10.6254 | -105.1794
10.6292 | -105.1842
10.6329 | -105.1841
10.6367 | -105.1851
10.6404 | -105.1828
10.6442 | -105.1799
10.6479 | -105.1769
10.6517 | -105.1702
10.6554 | -105.1609
10.6592 | -105.1514
10.6629 | -105.1402
10.6667 | -105.128
10.6704 | -105.1121
10.6742 | -105.0964
10.6779 | -105.0803
10.6817 | -105.0638
10.6854 | -105.0466
10.6892 | -105.0274
10.6929 | -105.0105
10.6967 | -104.9923
10.7004 | -104.9713
10.7042 | -104.949
10.7079 | -104.9288
10.7117 | -104.9118
10.7154 | -104.8959
10.7192 | -104.8819
10.7229 | -104.8739
10.7267 | -104.8701
10.7304 | -104.866
10.7342 | -104.8664
10.7379 | -104.871
10.7417 | -104.8794
10.7454 | -104.8933
10.7492 | -104.9105
10.7529 | -104.9318
10.7567 | -104.9548
10.7604 | -104.9752
10.7642 | -104.9991
10.7679 | -105.0199
10.7717 | -105.0435
10.7754 | -105.07
10.7792 | -105.0913
10.7829 | -105.1116
10.7867 | -105.1313
10.7904 | -105.1424
10.7942 | -105.1505
10.7979 | -105.1573
10.8017 | -105.1597
10.8054 | -105.1598
10.8092 | -105.1576
10.8129 | -105.1484
10.8167 | -105.1348
10.8204 | -105.1109
10.8242 | -105.098
10.8279 | -105.085
10.8317 | -105.0716
10.8354 | -105.0564
10.8392 | -105.041
10.8429 | -105.0294
10.8467 | -105.0213
10.8504 | -105.0093
10.8542 | -105.0004
10.8579 | -104.9944
10.8617 | -104.9918
10.8654 | -104.9934
10.8692 | -104.9959
10.8729 | -105.0027
10.8767 | -105.01
10.8804 | -105.0143
10.8842 | -105.0217
10.888 | -105.0345
10.8917 | -105.0492
10.8955 | -105.0652
10.8992 | -105.085
10.903 | -105.1028
10.9067 | -105.1198
10.9105 | -105.1396
10.9142 | -105.1601
10.918 | -105.1849
10.9217 | -105.2036
10.9255 | -105.2227
10.9292 | -105.2347
10.933 | -105.2448
10.9367 | -105.2534
10.9405 | -105.2581
10.9442 | -105.2623
10.948 | -105.2643
10.9517 | -105.2653
10.9555 | -105.2614
10.9592 | -105.2561
10.963 | -105.2429
10.9667 | -105.2313
10.9705 | -105.2237
10.9742 | -105.2123
10.978 | -105.2033
10.9817 | -105.1924
10.9855 | -105.1788
10.9892 | -105.16
10.993 | -105.1408
10.9967 | -105.1239
11.0005 | -105.1081
11.0042 | -105.0954
11.008 | -105.0858
11.0117 | -105.0799
11.0155 | -105.0718
11.0192 | -105.0678
11.023 | -105.0638
11.0267 | -105.0604
11.0305 | -105.0574
11.0342 | -105.0517
11.038 | -105.0531
11.0417 | -105.0611
11.0455 | -105.0652
11.0492 | -105.0679
11.053 | -105.071
11.0567 | -105.0708
11.0605 | -105.0664
11.0642 | -105.0607
11.068 | -105.059
11.0717 | -105.0564
11.0755 | -105.0514
11.0792 | -105.04
11.083 | -105.0295
11.0867 | -105.0197
11.0905 | -105.012
11.0942 | -105
11.098 | -104.9886
11.1017 | -104.9786
11.1055 | -104.9671
11.1092 | -104.956
11.113 | -104.944
11.1167 | -104.9306
11.1205 | -104.9229
11.1242 | -104.9226
11.128 | -104.9195
11.1317 | -104.9091
11.1355 | -104.9024
11.1392 | -104.8958
11.143 | -104.8916
11.1467 | -104.889
11.1505 | -104.8894
11.1542 | -104.8887
11.158 | -104.8885
11.1617 | -104.8916
11.1655 | -104.8937
11.1692 | -104.8956
11.173 | -104.9
11.1767 | -104.9103
11.1805 | -104.9233
11.1842 | -104.9374
11.188 | -104.9538
11.1917 | -104.9769
11.1955 | -105.0046
11.1992 | -105.0297
11.203 | -105.052
11.2068 | -105.0777
11.2105 | -105.1047
11.2143 | -105.1286
11.218 | -105.1561
11.2218 | -105.1824
11.2255 | -105.2083
11.2293 | -105.228
11.233 | -105.2449
11.2368 | -105.2612
11.2405 | -105.2755
11.2443 | -105.2885
11.248 | -105.2974
11.2518 | -105.305
11.2555 | -105.3142
11.2593 | -105.3169
11.263 | -105.3167
11.2668 | -105.3152
11.2705 | -105.3122
11.2743 | -105.3052
11.278 | -105.2991
11.2818 | -105.2859
11.2855 | -105.269
11.2893 | -105.2537
11.293 | -105.2447
11.2968 | -105.235
11.3005 | -105.2184
11.3043 | -105.2015
11.308 | -105.1898
11.3118 | -105.1787
11.3155 | -105.1694
11.3193 | -105.1609
11.323 | -105.1522
11.3268 | -105.1417
11.3305 | -105.1365
11.3343 | -105.132
11.338 | -105.1237
11.3418 | -105.1148
11.3455 | -105.1088
11.3493 | -105.1012
11.353 | -105.0938
11.3568 | -105.088
11.3605 | -105.0883
11.3643 | -105.0856
11.368 | -105.0825
11.3718 | -105.0824
11.3755 | -105.0787
11.3793 | -105.0759
11.383 | -105.0776
11.3868 | -105.0815
11.3905 | -105.0828
11.3943 | -105.0839
11.398 | -105.0882
11.4018 | -105.0945
11.4055 | -105.1024
11.4093 | -105.1103
11.413 | -105.1157
11.4168 | -105.1179
11.4205 | -105.1217
11.4243 | -105.1318
11.428 | -105.1382
11.4318 | -105.1433
11.4355 | -105.1485
11.4393 | -105.1546
11.443 | -105.1579
11.4468 | -105.1597
11.4505 | -105.1625
11.4543 | -105.1668
11.458 | -105.1699
11.4618 | -105.1728
11.4655 | -105.1799
11.4693 | -105.1833
11.473 | -105.1849
11.4768 | -105.1859
11.4805 | -105.1842
11.4843 | -105.1822
11.488 | -105.1803
11.4918 | -105.1814
11.4955 | -105.1796
11.4993 | -105.1796
11.503 | -105.183
11.5068 | -105.1846
11.5105 | -105.1859
11.5143 | -105.1879
11.5181 | -105.1896
11.5218 | -105.1936
11.5256 | -105.1994
11.5293 | -105.2013
11.5331 | -105.2015
11.5368 | -105.2064
11.5406 | -105.2083
11.5443 | -105.2119
11.5481 | -105.2157
11.5518 | -105.2179
11.5556 | -105.2206
11.5593 | -105.2239
11.5631 | -105.229
11.5668 | -105.2328
11.5706 | -105.2352
11.5743 | -105.2373
11.5781 | -105.2399
11.5818 | -105.2426
11.5856 | -105.2471
11.5893 | -105.2499
11.5931 | -105.2459
11.5968 | -105.2447
11.6006 | -105.2435
11.6043 | -105.2392
11.6081 | -105.2327
11.6118 | -105.2291
11.6156 | -105.2256
11.6193 | -105.2227
11.6231 | -105.2154
11.6268 | -105.2078
11.6306 | -105.2042
11.6343 | -105.2046
11.6381 | -105.2058
11.6418 | -105.212
11.6456 | -105.2207
11.6493 | -105.2284
11.6531 | -105.2359
11.6568 | -105.2445
11.6606 | -105.257
11.6643 | -105.2717
11.6681 | -105.2817
11.6718 | -105.2887
11.6756 | -105.2986
11.6793 | -105.31
11.6831 | -105.3231
11.6868 | -105.3333
11.6906 | -105.3418
11.6943 | -105.3482
11.6981 | -105.359
11.7018 | -105.3711
11.7056 | -105.3792
11.7093 | -105.3895
11.7131 | -105.399
11.7168 | -105.4071
11.7206 | -105.4105
11.7243 | -105.4148
11.7281 | -105.4192
11.7318 | -105.4176
11.7356 | -105.4133
11.7393 | -105.413
11.7431 | -105.4165
11.7468 | -105.416
11.7506 | -105.4167
11.7543 | -105.4138
11.7581 | -105.4099
11.7618 | -105.4078
11.7656 | -105.4048
11.7693 | -105.402
11.7731 | -105.4042
11.7768 | -105.4067
11.7806 | -105.407
11.7843 | -105.4117
11.7881 | -105.4157
11.7918 | -105.4235
11.7956 | -105.4346
11.7993 | -105.4463
11.8031 | -105.4572
11.8068 | -105.4696
11.8106 | -105.4794
11.8143 | -105.4937
11.8181 | -105.5131
11.8218 | -105.5322
11.8256 | -105.552
11.8293 | -105.5699
11.8331 | -105.5866
11.8369 | -105.6077
11.8406 | -105.6313
11.8444 | -105.6517
11.8481 | -105.6691
11.8519 | -105.6873
11.8556 | -105.7064
11.8594 | -105.7245
11.8631 | -105.7429
11.8669 | -105.7613
11.8706 | -105.7771
11.8744 | -105.7932
11.8781 | -105.8037
11.8819 | -105.8119
11.8856 | -105.8173
11.8894 | -105.822
11.8931 | -105.8248
11.8969 | -105.8273
11.9006 | -105.8228
11.9044 | -105.8247
11.9081 | -105.8236
11.9119 | -105.8154
11.9156 | -105.8034
11.9194 | -105.792
11.9231 | -105.7818
11.9269 | -105.7701
11.9306 | -105.7595
11.9344 | -105.7524
11.9381 | -105.7445
11.9419 | -105.7373
11.9456 | -105.7278
11.9494 | -105.7203
11.9531 | -105.7136
11.9569 | -105.7046
11.9606 | -105.6986
11.9644 | -105.6948
11.9681 | -105.6924
11.9719 | -105.6916
11.9756 | -105.6913
11.9794 | -105.6914
11.9831 | -105.6949
11.9869 | -105.6994
11.9906 | -105.702
11.9944 | -105.7038
11.9981 | -105.7024
12.0019 | -105.7006
12.0056 | -105.6997
12.0094 | -105.6981
12.0131 | -105.6926
12.0169 | -105.6884
12.0206 | -105.6811
12.0244 | -105.6719
12.0281 | -105.664
12.0319 | -105.6536
12.0356 | -105.643
12.0394 | -105.6259
12.0431 | -105.6117
12.0469 | -105.5965
12.0506 | -105.58
12.0544 | -105.5628
12.0581 | -105.5433
12.0619 | -105.5256
12.0656 | -105.5071
12.0694 | -105.4879
12.0731 | -105.47
12.0769 | -105.4527
12.0806 | -105.4345
12.0844 | -105.4129
12.0881 | -105.3978
12.0919 | -105.3838
12.0956 | -105.3736
12.0994 | -105.3611
12.1031 | -105.3512
12.1069 | -105.3388
12.1106 | -105.3248
12.1144 | -105.3145
12.1181 | -105.3043
12.1219 | -105.2954
12.1256 | -105.2891
12.1294 | -105.2848
12.1331 | -105.2835
12.1369 | -105.2815
12.1406 | -105.2791
12.1444 | -105.2746
12.1481 | -105.2695
12.1519 | -105.2724
12.1556 | -105.2747
12.1594 | -105.2773
12.1631 | -105.2797
12.1669 | -105.2876
12.1707 | -105.3
12.1744 | -105.3126
12.1782 | -105.3261
12.1819 | -105.3411
12.1857 | -105.3561
12.1894 | -105.3719
12.1932 | -105.3871
12.1969 | -105.4056
12.2007 | -105.4242
12.2044 | -105.4393
12.2082 | -105.4534
12.2119 | -105.4664
12.2157 | -105.482
12.2194 | -105.494
12.2232 | -105.5002
12.2269 | -105.5098
12.2307 | -105.5212
12.2344 | -105.5303
12.2382 | -105.5343
12.2419 | -105.5393
12.2457 | -105.5455
12.2494 | -105.5518
12.2532 | -105.5603
12.2569 | -105.5717
12.2607 | -105.5869
12.2644 | -105.5965
12.2682 | -105.6074
12.2719 | -105.6205
12.2757 | -105.6351
12.2794 | -105.6499
12.2832 | -105.6693
12.2869 | -105.6889
12.2907 | -105.707
12.2944 | -105.7274
12.2982 | -105.746
12.3019 | -105.7635
12.3057 | -105.7828
12.3094 | -105.8009
12.3132 | -105.8184
12.3169 | -105.8335
12.3207 | -105.8492
12.3244 | -105.868
12.3282 | -105.8782
12.3319 | -105.8849
12.3357 | -105.8886
12.3394 | -105.8938
12.3432 | -105.8947
12.3469 | -105.8938
12.3507 | -105.8935
12.3544 | -105.8879
12.3582 | -105.8789
12.3619 | -105.8704
12.3657 | -105.8628
12.3694 | -105.8494
12.3732 | -105.8305
12.3769 | -105.8149
12.3807 | -105.7997
12.3844 | -105.7809
12.3882 | -105.7646
12.3919 | -105.7492
12.3957 | -105.7329
12.3994 | -105.7233
12.4032 | -105.7131
12.4069 | -105.7059
12.4107 | -105.7044
12.4144 | -105.7031
12.4182 | -105.7031
12.4219 | -105.7049
12.4257 | -105.7083
12.4294 | -105.7131
12.4332 | -105.7183
12.4369 | -105.7255
12.4407 | -105.7341
12.4444 | -105.7394
12.4482 | -105.7434
12.4519 | -105.7453
12.4557 | -105.747
12.4594 | -105.748
12.4632 | -105.7473
12.4669 | -105.7434
12.4707 | -105.7376
12.4744 | -105.7333
12.4782 | -105.7262
12.4819 | -105.7172
12.4857 | -105.704
12.4895 | -105.6883
12.4932 | -105.6712
12.497 | -105.6545
12.5007 | -105.6357
12.5045 | -105.6139
12.5082 | -105.5952
12.512 | -105.5807
12.5157 | -105.565
12.5195 | -105.5444
12.5232 | -105.5232
12.527 | -105.5052
12.5307 | -105.489
12.5345 | -105.4775
12.5382 | -105.4685
12.542 | -105.4599
12.5457 | -105.4545
12.5495 | -105.4546
12.5532 | -105.4598
12.557 | -105.465
12.5607 | -105.4726
12.5645 | -105.4858
12.5682 | -105.4966
12.572 | -105.5061
12.5757 | -105.5169
12.5795 | -105.5276
12.5832 | -105.5384
12.587 | -105.5477
12.5907 | -105.5561
12.5945 | -105.5634
12.5982 | -105.571
12.602 | -105.576
12.6057 | -105.5779
12.6095 | -105.5806
12.6132 | -105.5817
12.617 | -105.5774
12.6207 | -105.573
12.6245 | -105.571
12.6282 | -105.5638
12.632 | -105.5559
12.6357 | -105.5477
12.6395 | -105.5384
12.6432 | -105.5287
12.647 | -105.5181
12.6507 | -105.511
12.6545 | -105.5038
12.6582 | -105.491
12.662 | -105.4834
12.6657 | -105.4775
12.6695 | -105.4763
12.6732 | -105.4786
12.677 | -105.4827
12.6807 | -105.4876
12.6845 | -105.4957
12.6882 | -105.5045
12.692 | -105.5167
12.6957 | -105.5303
12.6995 | -105.5449
12.7032 | -105.5626
12.707 | -105.5851
12.7107 | -105.6047
12.7145 | -105.6266
12.7182 | -105.6458
12.722 | -105.6632
12.7257 | -105.6793
12.7295 | -105.6926
12.7332 | -105.7037
12.737 | -105.7118
12.7407 | -105.7168
12.7445 | -105.7193
12.7482 | -105.7201
12.752 | -105.7207
12.7557 | -105.7191
12.7595 | -105.7146
12.7632 | -105.705
12.767 | -105.693
12.7707 | -105.6824
12.7745 | -105.6704
12.7782 | -105.6555
12.782 | -105.6381
12.7857 | -105.6224
12.7895 | -105.6079
12.7932 | -105.5902
12.797 | -105.5688
12.8008 | -105.5505
12.8045 | -105.5312
12.8083 | -105.5083
12.812 | -105.4922
12.8158 | -105.4775
12.8195 | -105.4617
12.8233 | -105.4483
12.827 | -105.4346
12.8308 | -105.423
12.8345 | -105.412
12.8383 | -105.4002
12.842 | -105.3909
12.8458 | -105.3866
12.8495 | -105.3817
12.8533 | -105.3755
12.857 | -105.3733
12.8608 | -105.3738
12.8645 | -105.3765
12.8683 | -105.3801
12.872 | -105.3871
12.8758 | -105.3927
12.8795 | -105.4025
12.8833 | -105.4158
12.887 | -105.4315
12.8908 | -105.4453
12.8945 | -105.4629
12.8983 | -105.4814
12.902 | -105.501
12.9058 | -105.5184
12.9095 | -105.5382
12.9133 | -105.5596
12.917 | -105.5804
12.9208 | -105.5992
12.9245 | -105.6172
12.9283 | -105.6355
12.932 | -105.651
12.9358 | -105.6614
12.9395 | -105.6701
12.9433 | -105.6769
12.947 | -105.6817
12.9508 | -105.684
12.9545 | -105.6812
12.9583 | -105.6763
12.962 | -105.6737
12.9658 | -105.6702
12.9695 | -105.6659
12.9733 | -105.6631
12.977 | -105.6603
12.9808 | -105.6574
12.9845 | -105.6586
12.9883 | -105.6603
12.992 | -105.661
12.9958 | -105.6656
12.9995 | -105.6711
13.0033 | -105.6734
13.007 | -105.677
13.0108 | -105.6856
13.0145 | -105.6969
13.0183 | -105.7066
13.022 | -105.7171
13.0258 | -105.7328
13.0295 | -105.7467
13.0333 | -105.7603
13.037 | -105.7743
13.0408 | -105.7872
13.0445 | -105.799
13.0483 | -105.8097
13.052 | -105.8194
13.0558 | -105.8275
13.0595 | -105.8318
13.0633 | -105.8399
13.067 | -105.8462
13.0708 | -105.8471
13.0745 | -105.8492
13.0783 | -105.8523
13.082 | -105.854
13.0858 | -105.8551
13.0895 | -105.8548
13.0933 | -105.856
13.097 | -105.8556
13.1008 | -105.8529
13.1045 | -105.851
13.1083 | -105.8502
13.112 | -105.8476
13.1158 | -105.8433
13.1196 | -105.8366
13.1233 | -105.8326
13.1271 | -105.8277
13.1308 | -105.8226
13.1346 | -105.8154
13.1383 | -105.8057
13.1421 | -105.7969
13.1458 | -105.7919
13.1496 | -105.7846
13.1533 | -105.7739
13.1571 | -105.7633
13.1608 | -105.7505
13.1646 | -105.7411
13.1683 | -105.7354
13.1721 | -105.7274
13.1758 | -105.718
13.1796 | -105.7076
13.1833 | -105.698
13.1871 | -105.6877
13.1908 | -105.6759
13.1946 | -105.664
13.1983 | -105.654
13.2021 | -105.6425
13.2058 | -105.6279
13.2096 | -105.6146
13.2133 | -105.6025
13.2171 | -105.5854
13.2208 | -105.5698
13.2246 | -105.5503
13.2283 | -105.5359
13.2321 | -105.5228
13.2358 | -105.5064
13.2396 | -105.4877
13.2433 | -105.4685
13.2471 | -105.4528
13.2508 | -105.4373
13.2546 | -105.418
13.2583 | -105.3993
13.2621 | -105.385
13.2658 | -105.3689
13.2696 | -105.3543
13.2733 | -105.3409
13.2771 | -105.3291
13.2808 | -105.3199
13.2846 | -105.3139
13.2883 | -105.309
13.2921 | -105.3044
13.2958 | -105.3023
13.2996 | -105.3023
13.3033 | -105.3092
13.3071 | -105.3194
13.3108 | -105.3322
13.3146 | -105.3446
13.3183 | -105.3579
13.3221 | -105.3698
13.3258 | -105.385
13.3296 | -105.4037
13.3333 | -105.4199
13.3371 | -105.4332
13.3408 | -105.4496
13.3446 | -105.4641
13.3483 | -105.4783
13.3521 | -105.4963
13.3558 | -105.5115
13.3596 | -105.5248
13.3633 | -105.5359
13.3671 | -105.5505
13.3708 | -105.5625
13.3746 | -105.5766
13.3783 | -105.59
13.3821 | -105.6053
13.3858 | -105.6245
13.3896 | -105.641
13.3933 | -105.6518
13.3971 | -105.6663
13.4008 | -105.6797
13.4046 | -105.6927
13.4083 | -105.7079
13.4121 | -105.7212
13.4158 | -105.7335
13.4196 | -105.741
13.4233 | -105.7502
13.4271 | -105.7616
13.4308 | -105.7747
13.4346 | -105.7904
13.4383 | -105.8038
13.4421 | -105.8153
13.4459 | -105.824
13.4496 | -105.8336
13.4534 | -105.8386
13.4571 | -105.8378
13.4609 | -105.836
13.4646 | -105.8372
13.4684 | -105.8372
13.4721 | -105.8364
13.4759 | -105.8343
13.4796 | -105.8288
13.4834 | -105.8194
13.4871 | -105.8066
13.4909 | -105.7896
13.4946 | -105.7685
13.4984 | -105.7448
13.5021 | -105.721
13.5059 | -105.6979
13.5096 | -105.676
13.5134 | -105.6517
13.5171 | -105.6244
13.5209 | -105.5926
13.5246 | -105.5671
13.5284 | -105.5406
13.5321 | -105.5148
13.5359 | -105.4888
13.5396 | -105.4651
13.5434 | -105.4415
13.5471 | -105.4191
13.5509 | -105.4005
13.5546 | -105.3863
13.5584 | -105.3762
13.5621 | -105.3639
13.5659 | -105.3539
13.5696 | -105.3448
13.5734 | -105.3413
13.5771 | -105.3392
13.5809 | -105.338
13.5846 | -105.3368
13.5884 | -105.3368
13.5921 | -105.3392
13.5959 | -105.3428
13.5996 | -105.3491
13.6034 | -105.3546
13.6071 | -105.3592
13.6109 | -105.3648
13.6146 | -105.3717
13.6184 | -105.3778
13.6221 | -105.3833
13.6259 | -105.3905
13.6296 | -105.3959
13.6334 | -105.4019
13.6371 | -105.409
13.6409 | -105.418
13.6446 | -105.4268
13.6484 | -105.4398
13.6521 | -105.4572
13.6559 | -105.4719
13.6596 | -105.4818
13.6634 | -105.4933
13.6671 | -105.5043
13.6709 | -105.5178
13.6746 | -105.5337
13.6784 | -105.555
13.6821 | -105.5783
13.6859 | -105.602
13.6896 | -105.6236
13.6934 | -105.6474
13.6971 | -105.6708
13.7009 | -105.6932
13.7046 | -105.7145
13.7084 | -105.7383
13.7121 | -105.7638
13.7159 | -105.7918
13.7196 | -105.8161
13.7234 | -105.8374
13.7271 | -105.8571
13.7309 | -105.8737
13.7346 | -105.8868
13.7384 | -105.8991
13.7421 | -105.9052
13.7459 | -105.9071
13.7496 | -105.9046
13.7534 | -105.8993
13.7571 | -105.8902
13.7609 | -105.8753
13.7647 | -105.8602
13.7684 | -105.8376
13.7722 | -105.8071
13.7759 | -105.7773
13.7797 | -105.7445
13.7834 | -105.7111
13.7872 | -105.6784
13.7909 | -105.6417
13.7947 | -105.6023
13.7984 | -105.5594
13.8022 | -105.5221
13.8059 | -105.4858
13.8097 | -105.448
13.8134 | -105.4121
13.8172 | -105.3777
13.8209 | -105.3451
13.8247 | -105.3109
13.8284 | -105.2795
13.8322 | -105.2547
13.8359 | -105.2325
13.8397 | -105.2138
13.8434 | -105.1989
13.8472 | -105.1863
13.8509 | -105.1792
13.8547 | -105.1753
13.8584 | -105.1749
13.8622 | -105.1754
13.8659 | -105.181
13.8697 | -105.1879
13.8734 | -105.1954
13.8772 | -105.2049
13.8809 | -105.2155
13.8847 | -105.2295
13.8884 | -105.2438
13.8922 | -105.2565
13.8959 | -105.2698
13.8997 | -105.282
13.9034 | -105.2958
13.9072 | -105.3083
13.9109 | -105.3179
13.9147 | -105.3254
13.9184 | -105.3333
13.9222 | -105.3414
13.9259 | -105.3496
13.9297 | -105.3574
13.9334 | -105.3645
13.9372 | -105.3698
13.9409 | -105.3771
13.9447 | -105.3809
13.9484 | -105.3836
13.9522 | -105.3844
13.9559 | -105.3867
13.9597 | -105.3924
13.9634 | -105.3927
13.9672 | -105.3942
13.9709 | -105.3952
13.9747 | -105.3922
13.9784 | -105.3903
13.9822 | -105.3882
13.9859 | -105.3831
13.9897 | -105.3758
13.9934 | -105.3674
13.9972 | -105.3561
14.0009 | -105.3392
14.0047 | -105.3236
14.0084 | -105.3065
14.0122 | -105.2861
14.0159 | -105.2663
14.0197 | -105.2453
14.0234 | -105.2231
14.0272 | -105.1966
14.0309 | -105.169
14.0347 | -105.1457
14.0384 | -105.1236
14.0422 | -105.101
14.0459 | -105.0793
14.0497 | -105.0577
14.0534 | -105.0399
14.0572 | -105.0253
14.0609 | -105.0146
14.0647 | -105.0239
14.0684 | -105.0197
14.0722 | -105.0165
14.0759 | -105.0207
14.0797 | -105.0274
14.0835 | -105.0373
14.0872 | -105.0493
14.091 | -105.0637
14.0947 | -105.0813
14.0985 | -105.1024
14.1022 | -105.1247
14.106 | -105.146
14.1097 | -105.1718
14.1135 | -105.1996
14.1172 | -105.228
14.121 | -105.2558
14.1247 | -105.2844
14.1285 | -105.3131
14.1322 | -105.3422
14.136 | -105.3662
14.1397 | -105.3906
14.1435 | -105.4149
14.1472 | -105.4405
14.151 | -105.4643
14.1547 | -105.4864
14.1585 | -105.5041
14.1622 | -105.5216
14.166 | -105.5386
14.1697 | -105.5535
14.1735 | -105.5653
14.1772 | -105.5746
14.181 | -105.5799
14.1847 | -105.582
14.1885 | -105.5824
14.1922 | -105.5805
14.196 | -105.5787
14.1997 | -105.5749
14.2035 | -105.5699
14.2072 | -105.563
14.211 | -105.5548
14.2147 | -105.549
14.2185 | -105.5427
14.2222 | -105.5368
14.226 | -105.5318
14.2297 | -105.5283
14.2335 | -105.5203
14.2372 | -105.5112
14.241 | -105.5023
14.2447 | -105.4926
14.2485 | -105.4858
14.2522 | -105.4816
14.256 | -105.4785
14.2597 | -105.4745
14.2635 | -105.4716
14.2672 | -105.4697
14.271 | -105.4675
14.2747 | -105.467
14.2785 | -105.4695
14.2822 | -105.4714
14.286 | -105.4722
14.2897 | -105.4733
14.2935 | -105.4763
14.2972 | -105.4805
14.301 | -105.4867
14.3047 | -105.4912
14.3085 | -105.4962
14.3122 | -105.505
14.316 | -105.5125
14.3197 | -105.5209
14.3235 | -105.5313
14.3272 | -105.5396
14.331 | -105.5478
14.3347 | -105.5594
14.3385 | -105.5701
14.3422 | -105.5837
14.346 | -105.5922
14.3497 | -105.5975
14.3535 | -105.6011
14.3572 | -105.607
14.361 | -105.613
14.3647 | -105.6189
14.3685 | -105.6219
14.3722 | -105.6239
14.376 | -105.6244
14.3797 | -105.6253
14.3835 | -105.6248
14.3872 | -105.6235
14.391 | -105.6217
14.3947 | -105.6212
14.3985 | -105.6208
14.4023 | -105.6186
14.406 | -105.6184
14.4098 | -105.617
14.4135 | -105.6132
14.4173 | -105.6104
14.421 | -105.6054
14.4248 | -105.6026
14.4285 | -105.6005
14.4323 | -105.5996
14.436 | -105.5989
14.4398 | -105.5997
14.4435 | -105.601
14.4473 | -105.6028
14.451 | -105.6057
14.4548 | -105.6081
14.4585 | -105.6118
14.4623 | -105.6168
14.466 | -105.6236
14.4698 | -105.6306
14.4735 | -105.638
14.4773 | -105.6478
14.481 | -105.6583
14.4848 | -105.6711
14.4885 | -105.6868
14.4923 | -105.7042
14.496 | -105.723
14.4998 | -105.7439
14.5035 | -105.7634
14.5073 | -105.7798
14.511 | -105.7981
14.5148 | -105.8194
14.5185 | -105.8369
14.5223 | -105.8504
14.526 | -105.8663
14.5298 | -105.8805
14.5335 | -105.898
14.5373 | -105.9121
14.541 | -105.9221
14.5448 | -105.9303
14.5485 | -105.9343
14.5523 | -105.9418
14.556 | -105.9446
14.5598 | -105.942
14.5635 | -105.9359
14.5673 | -105.9262
14.571 | -105.9127
14.5748 | -105.8982
14.5785 | -105.8834
14.5823 | -105.8682
14.586 | -105.8521
14.5898 | -105.8367
14.5935 | -105.8199
14.5973 | -105.8004
14.601 | -105.782
14.6048 | -105.7637
14.6085 | -105.7502
14.6123 | -105.7355
14.616 | -105.721
14.6198 | -105.7041
14.6235 | -105.6904
14.6273 | -105.6802
14.631 | -105.6704
14.6348 | -105.661
14.6385 | -105.6546
14.6423 | -105.6501
14.646 | -105.647
14.6498 | -105.6466
14.6535 | -105.6479
14.6573 | -105.6513
14.661 | -105.6546
14.6648 | -105.6598
14.6685 | -105.6652
14.6723 | -105.6715
14.676 | -105.6805
14.6798 | -105.689
14.6835 | -105.6993
14.6873 | -105.7112
14.691 | -105.7259
14.6948 | -105.7386
14.6985 | -105.751
14.7023 | -105.7664
14.706 | -105.7803
14.7098 | -105.7941
14.7135 | -105.8105
14.7173 | -105.8283
14.721 | -105.8437
14.7248 | -105.8578
14.7286 | -105.8736
14.7323 | -105.891
14.7361 | -105.9076
14.7398 | -105.921
14.7436 | -105.9329
14.7473 | -105.9465
14.7511 | -105.9598
14.7548 | -105.9723
14.7586 | -105.9875
14.7623 | -105.9977
14.7661 | -106.0078
14.7698 | -106.019
14.7736 | -106.0284
14.7773 | -106.0342
14.7811 | -106.0397
14.7848 | -106.0442
14.7886 | -106.0455
14.7923 | -106.0458
14.7961 | -106.0448
14.7998 | -106.0414
14.8036 | -106.0374
14.8073 | -106.0357
14.8111 | -106.0294
14.8148 | -106.0254
14.8186 | -106.0231
14.8223 | -106.0213
14.8261 | -106.0173
14.8298 | -106.0116
14.8336 | -106.0127
14.8373 | -106.0113
14.8411 | -106.0064
14.8448 | -106.002
14.8486 | -106.0013
14.8523 | -106.001
14.8561 | -106.0013
14.8598 | -106.005
14.8636 | -106.0082
14.8673 | -106.0126
14.8711 | -106.017
14.8748 | -106.024
14.8786 | -106.0307
14.8823 | -106.0395
14.8861 | -106.0486
14.8898 | -106.0565
14.8936 | -106.0634
14.8973 | -106.0693
14.9011 | -106.0761
14.9048 | -106.0844
14.9086 | -106.0916
14.9123 | -106.0961
14.9161 | -106.0998
14.9198 | -106.1048
14.9236 | -106.1059
14.9273 | -106.1064
14.9311 | -106.105
14.9348 | -106.1026
14.9386 | -106.0972
14.9423 | -106.0904
14.9461 | -106.084
14.9498 | -106.0719
14.9536 | -106.06
14.9573 | -106.044
14.9611 | -106.0242
14.9648 | -106.0047
14.9686 | -105.9856
14.9723 | -105.9663
14.9761 | -105.9464
14.9798 | -105.9291
14.9836 | -105.9086
14.9873 | -105.8895
14.9911 | -105.8737
14.9948 | -105.859
14.9986 | -105.8446
15.0023 | -105.8317
15.0061 | -105.8182
15.0098 | -105.8072
15.0136 | -105.7988
15.0173 | -105.7938
15.0211 | -105.7869
15.0248 | -105.7825
15.0286 | -105.7797
15.0323 | -105.7788
15.0361 | -105.7809
15.0398 | -105.7852
15.0436 | -105.7916
15.0474 | -105.7951
15.0511 | -105.8025
15.0549 | -105.8138
15.0586 | -105.8239
15.0624 | -105.8317
15.0661 | -105.8416
15.0699 | -105.8519
15.0736 | -105.8636
15.0774 | -105.8768
15.0811 | -105.8883
15.0849 | -105.9022
15.0886 | -105.9139
15.0924 | -105.9273
15.0961 | -105.9356
15.0999 | -105.9451
15.1036 | -105.9546
15.1074 | -105.9616
15.1111 | -105.9643
15.1149 | -105.9668
15.1186 | -105.9662
15.1224 | -105.9663
15.1261 | -105.9648
15.1299 | -105.9624
15.1336 | -105.9558
15.1374 | -105.9497
15.1411 | -105.94
15.1449 | -105.9302
15.1486 | -105.9159
15.1524 | -105.9059
15.1561 | -105.8945
15.1599 | -105.8804
15.1636 | -105.8662
15.1674 | -105.8553
15.1711 | -105.8487
15.1749 | -105.8418
15.1786 | -105.8315
15.1824 | -105.822
15.1861 | -105.8131
15.1899 | -105.8036
15.1936 | -105.7972
15.1974 | -105.7908
15.2011 | -105.7889
15.2049 | -105.7856
15.2086 | -105.7831
15.2124 | -105.7813
15.2161 | -105.7801
15.2199 | -105.7803
15.2236 | -105.7834
15.2274 | -105.7865
15.2311 | -105.7907
15.2349 | -105.7999
15.2386 | -105.813
15.2424 | -105.8241
15.2461 | -105.8344
15.2499 | -105.845
15.2536 | -105.8573
15.2574 | -105.8727
15.2611 | -105.8901
15.2649 | -105.9095
15.2686 | -105.9284
15.2724 | -105.9473
15.2761 | -105.9669
15.2799 | -105.9855
15.2836 | -106.0055
15.2874 | -106.0257
15.2911 | -106.0417
15.2949 | -106.0578
15.2986 | -106.0762
15.3024 | -106.0939
15.3061 | -106.113
15.3099 | -106.1289
15.3136 | -106.1456
15.3174 | -106.1601
15.3211 | -106.1742
15.3249 | -106.1903
15.3286 | -106.2037
15.3324 | -106.2118
15.3361 | -106.2189
15.3399 | -106.2224
15.3436 | -106.2295
15.3474 | -106.2333
15.3511 | -106.2373
15.3549 | -106.2385
15.3587 | -106.2405
15.3624 | -106.2412
15.3662 | -106.2393
15.3699 | -106.2352
15.3737 | -106.2275
15.3774 | -106.2184
15.3812 | -106.2116
15.3849 | -106.2038
15.3887 | -106.1927
15.3924 | -106.1799
15.3962 | -106.169
15.3999 | -106.1593
15.4037 | -106.1462
15.4074 | -106.1301
15.4112 | -106.1138
15.4149 | -106.0969
15.4187 | -106.0795
15.4224 | -106.0611
15.4262 | -106.0451
15.4299 | -106.0321
15.4337 | -106.0198
15.4374 | -106.0088
15.4412 | -106.0016
15.4449 | -105.9946
15.4487 | -105.9883
15.4524 | -105.9829
15.4562 | -105.9736
15.4599 | -105.9672
15.4637 | -105.964
15.4674 | -105.9651
15.4712 | -105.9664
15.4749 | -105.9679
15.4787 | -105.9711
15.4824 | -105.9752
15.4862 | -105.9804
15.4899 | -105.9821
15.4937 | -105.9833
15.4974 | -105.9893
15.5012 | -105.9966
15.5049 | -106.0022
15.5087 | -106.0095
15.5124 | -106.0179
15.5162 | -106.0251
15.5199 | -106.0333
15.5237 | -106.0417
15.5274 | -106.0494
15.5312 | -106.0598
15.5349 | -106.0714
15.5387 | -106.0801
15.5424 | -106.0895
15.5462 | -106.0985
15.5499 | -106.111
15.5537 | -106.1228
15.5574 | -106.1339
15.5612 | -106.1445
15.5649 | -106.1511
15.5687 | -106.1592
15.5724 | -106.1663
15.5762 | -106.1745
15.5799 | -106.1775
15.5837 | -106.1748
15.5874 | -106.1746
15.5912 | -106.1738
15.5949 | -106.1691
15.5987 | -106.1624
15.6024 | -106.1512
15.6062 | -106.1384
15.6099 | -106.1212
15.6137 | -106.1011
15.6174 | -106.0762
15.6212 | -106.053
15.6249 | -106.025
15.6287 | -105.996
15.6324 | -105.9675
15.6362 | -105.9347
15.6399 | -105.8995
15.6437 | -105.8615
15.6474 | -105.8217
15.6512 | -105.7792
15.6549 | -105.7385
15.6587 | -105.6973
15.6624 | -105.6547
15.6662 | -105.6129
15.6699 | -105.5702
15.6737 | -105.5308
15.6775 | -105.4964
15.6812 | -105.462
15.685 | -105.4283
15.6887 | -105.4071
15.6925 | -105.3747
15.6962 | -105.3461
15.7 | -105.3235
15.7037 | -105.3016
15.7075 | -105.2827
15.7112 | -105.2675
15.715 | -105.2559
15.7187 | -105.2461
15.7225 | -105.2375
15.7262 | -105.2304
15.73 | -105.2268
15.7337 | -105.2258
15.7375 | -105.2268
15.7412 | -105.23
15.745 | -105.2329
15.7487 | -105.2387
15.7525 | -105.2477
15.7562 | -105.2558
15.76 | -105.2676
15.7637 | -105.2838
15.7675 | -105.2999
15.7712 | -105.3161
15.775 | -105.3352
15.7787 | -105.3578
15.7825 | -105.3799
15.7862 | -105.4035
15.79 | -105.4269
15.7937 | -105.4503
15.7975 | -105.4748
15.8012 | -105.4983
15.805 | -105.5221
15.8087 | -105.5497
15.8125 | -105.5712
15.8162 | -105.5911
15.82 | -105.608
15.8237 | -105.6242
15.8275 | -105.6372
15.8312 | -105.6441
15.835 | -105.6509
15.8387 | -105.6529
15.8425 | -105.6545
15.8462 | -105.6486
15.85 | -105.639
15.8537 | -105.631
15.8575 | -105.6165
15.8612 | -105.602
15.865 | -105.5819
15.8687 | -105.5607
15.8725 | -105.5372
15.8762 | -105.5092
15.88 | -105.4779
15.8837 | -105.4457
15.8875 | -105.4119
15.8912 | -105.3784
15.895 | -105.3445
15.8987 | -105.3073
15.9025 | -105.2675
15.9062 | -105.2303
15.91 | -105.1951
15.9137 | -105.158
15.9175 | -105.1233
15.9212 | -105.0904
15.925 | -105.057
15.9287 | -105.0212
15.9325 | -104.9914
15.9362 | -104.9653
15.94 | -104.9396
15.9437 | -104.9177
15.9475 | -104.8985
15.9512 | -104.8783
15.955 | -104.8629
15.9587 | -104.8508
15.9625 | -104.8383
15.9662 | -104.8282
15.97 | -104.8211
15.9737 | -104.8155
15.9775 | -104.8122
15.9812 | -104.8107
15.985 | -104.8108
15.9887 | -104.8119
15.9925 | -104.8128
15.9962 | -104.8162
16 | -104.819
16.0038 | -104.8182
16.0075 | -104.8185
16.0113 | -104.822
16.015 | -104.8262
16.0188 | -104.8299
16.0225 | -104.8338
16.0263 | -104.8363
16.03 | -104.8378
16.0338 | -104.8405
16.0375 | -104.8411
16.0413 | -104.8431
16.045 | -104.8454
16.0488 | -104.8504
16.0525 | -104.8516
16.0563 | -104.8532
16.06 | -104.8554
16.0638 | -104.8566
16.0675 | -104.8587
16.0713 | -104.8587
16.075 | -104.8584
16.0788 | -104.8623
16.0825 | -104.8613
16.0863 | -104.8614
16.09 | -104.8623
16.0938 | -104.8657
16.0975 | -104.8712
16.1013 | -104.8749
16.105 | -104.8803
16.1088 | -104.8877
16.1125 | -104.8945
16.1163 | -104.9028
16.12 | -104.9112
16.1238 | -104.921
16.1275 | -104.931
16.1313 | -104.9431
16.135 | -104.9526
16.1388 | -104.9635
16.1425 | -104.9785
16.1463 | -104.9924
16.15 | -105.0069
16.1538 | -105.0244
16.1575 | -105.0413
16.1613 | -105.0549
16.165 | -105.0694
16.1688 | -105.0852
16.1725 | -105.1017
16.1763 | -105.1166
16.18 | -105.1339
16.1838 | -105.148
16.1875 | -105.1588
16.1913 | -105.1709
16.195 | -105.1819
16.1988 | -105.1923
16.2025 | -105.2007
16.2063 | -105.2045
16.21 | -105.2055
16.2138 | -105.207
16.2175 | -105.2054
16.2213 | -105.1987
16.225 | -105.1865
16.2288 | -105.1772
16.2325 | -105.152
16.2363 | -105.1378
16.24 | -105.1192
16.2438 | -105.0954
16.2475 | -105.0683
16.2513 | -105.0425
16.255 | -105.0134
16.2588 | -104.9845
16.2625 | -104.9501
16.2663 | -104.9111
16.27 | -104.8742
16.2738 | -104.8314
16.2775 | -104.7873
16.2813 | -104.7437
16.285 | -104.6994
16.2888 | -104.6534
16.2925 | -104.6079
16.2963 | -104.563
16.3 | -104.5153
16.3038 | -104.47
16.3075 | -104.4271
16.3113 | -104.3816
16.315 | -104.3365
16.3188 | -104.2921
16.3225 | -104.2488
16.3263 | -104.2095
16.3301 | -104.169
16.3338 | -104.133
16.3376 | -104.0989
16.3413 | -104.0672
16.3451 | -104.0401
16.3488 | -104.0148
16.3526 | -103.9932
16.3563 | -103.976
16.3601 | -103.9645
16.3638 | -103.956
16.3676 | -103.9537
16.3713 | -103.9547
16.3751 | -103.9585
16.3788 | -103.9659
16.3826 | -103.9799
16.3863 | -103.9988
16.3901 | -104.0202
16.3938 | -104.0465
16.3976 | -104.0768
16.4013 | -104.1122
16.4051 | -104.1502
16.4088 | -104.1914
16.4126 | -104.235
16.4163 | -104.2806
16.4201 | -104.3273
16.4238 | -104.3732
16.4276 | -104.4203
16.4313 | -104.465
16.4351 | -104.5108
16.4388 | -104.558
16.4426 | -104.6045
16.4463 | -104.6468
16.4501 | -104.6862
16.4538 | -104.7263
16.4576 | -104.7634
16.4613 | -104.7945
16.4651 | -104.8225
16.4688 | -104.8493
16.4726 | -104.8733
16.4763 | -104.8944
16.4801 | -104.913
16.4838 | -104.9291
16.4876 | -104.938
16.4913 | -104.9461
16.4951 | -104.9523
16.4988 | -104.9589
16.5026 | -104.9638
16.5063 | -104.9618
16.5101 | -104.9605
16.5138 | -104.9533
16.5176 | -104.9419
16.5213 | -104.9294
16.5251 | -104.9159
16.5288 | -104.9003
16.5326 | -104.8811
16.5363 | -104.8675
16.5401 | -104.8524
16.5438 | -104.836
16.5476 | -104.821
16.5513 | -104.8079
16.5551 | -104.7954
16.5588 | -104.7839
16.5626 | -104.7708
16.5663 | -104.7612
16.5701 | -104.7531
16.5738 | -104.7463
16.5776 | -104.7415
16.5813 | -104.7362
16.5851 | -104.7359
16.5888 | -104.7398
16.5926 | -104.7416
16.5963 | -104.7416
16.6001 | -104.7462
16.6038 | -104.7539
16.6076 | -104.7627
16.6113 | -104.7717
16.6151 | -104.7853
16.6188 | -104.796
16.6226 | -104.8056
16.6263 | -104.8174
16.6301 | -104.8284
16.6338 | -104.8393
16.6376 | -104.8499
16.6413 | -104.8614
16.6451 | -104.8721
16.6489 | -104.882
16.6526 | -104.894
16.6564 | -104.9063
16.6601 | -104.9188
16.6639 | -104.9285
16.6676 | -104.9379
16.6714 | -104.9472
16.6751 | -104.9562
16.6789 | -104.9643
16.6826 | -104.9755
16.6864 | -104.9869
16.6901 | -104.9962
16.6939 | -105.0075
16.6976 | -105.0174
16.7014 | -105.0276
16.7051 | -105.0367
16.7089 | -105.0458
16.7126 | -105.0559
16.7164 | -105.0647
16.7201 | -105.074
16.7239 | -105.0844
16.7276 | -105.0931
16.7314 | -105.1003
16.7351 | -105.109
16.7389 | -105.1192
16.7426 | -105.1289
16.7464 | -105.1379
16.7501 | -105.148
16.7539 | -105.1555
16.7576 | -105.164
16.7614 | -105.1725
16.7651 | -105.1818
16.7689 | -105.1918
16.7726 | -105.1972
16.7764 | -105.1991
16.7801 | -105.2042
16.7839 | -105.2094
16.7876 | -105.2159
16.7914 | -105.2217
16.7951 | -105.2257
16.7989 | -105.228
16.8026 | -105.232
16.8064 | -105.2377
16.8101 | -105.2393
16.8139 | -105.242
16.8176 | -105.2447
16.8214 | -105.2489
16.8251 | -105.2501
16.8289 | -105.2475
16.8326 | -105.2476
16.8364 | -105.2478
16.8401 | -105.2448
16.8439 | -105.2401
16.8476 | -105.2377
16.8514 | -105.2369
16.8551 | -105.2351
16.8589 | -105.2352
16.8626 | -105.2339
16.8664 | -105.2353
16.8701 | -105.2367
16.8739 | -105.238
16.8776 | -105.237
16.8814 | -105.2366
16.8851 | -105.2403
16.8889 | -105.244
16.8926 | -105.246
16.8964 | -105.2492
16.9001 | -105.2521
16.9039 | -105.2574
16.9076 | -105.2637
16.9114 | -105.2693
16.9151 | -105.2742
16.9189 | -105.2797
16.9226 | -105.2831
16.9264 | -105.2884
16.9301 | -105.2953
16.9339 | -105.3017
16.9376 | -105.3106
16.9414 | -105.3185
16.9451 | -105.3264
16.9489 | -105.3349
16.9526 | -105.3405
16.9564 | -105.3464
16.9601 | -105.3527
16.9639 | -105.359
16.9677 | -105.3652
16.9714 | -105.3709
16.9752 | -105.3749
16.9789 | -105.3793
16.9827 | -105.3851
16.9864 | -105.392
16.9902 | -105.3986
16.9939 | -105.4064
16.9977 | -105.4105
17.0014 | -105.4144
17.0052 | -105.4181
17.0089 | -105.4234
17.0127 | -105.4278
17.0164 | -105.4344
17.0202 | -105.444
17.0239 | -105.4514
17.0277 | -105.4575
17.0314 | -105.4647
17.0352 | -105.4741
17.0389 | -105.4857
17.0427 | -105.4984
17.0464 | -105.5078
17.0502 | -105.5163
17.0539 | -105.5308
17.0577 | -105.5428
17.0614 | -105.5501
17.0652 | -105.5572
17.0689 | -105.5637
17.0727 | -105.5702
17.0764 | -105.5749
17.0802 | -105.5767
17.0839 | -105.5767
17.0877 | -105.5726
17.0914 | -105.5674
17.0952 | -105.5621
17.0989 | -105.554
17.1027 | -105.544
17.1064 | -105.5316
17.1102 | -105.5125
17.1139 | -105.4929
17.1177 | -105.4721
17.1214 | -105.4512
17.1252 | -105.4264
17.1289 | -105.3981
17.1327 | -105.3668
17.1364 | -105.3347
17.1402 | -105.3017
17.1439 | -105.2685
17.1477 | -105.2337
17.1514 | -105.1985
17.1552 | -105.1609
17.1589 | -105.1197
17.1627 | -105.0803
17.1664 | -105.0408
17.1702 | -105.0011
17.1739 | -104.9641
17.1777 | -104.9261
17.1814 | -104.8868
17.1852 | -104.8482
17.1889 | -104.8091
17.1927 | -104.7735
17.1964 | -104.7407
17.2002 | -104.7094
17.2039 | -104.6782
17.2077 | -104.6484
17.2114 | -104.6196
17.2152 | -104.59
17.2189 | -104.5624
17.2227 | -104.537
17.2264 | -104.5159
17.2302 | -104.4973
17.2339 | -104.479
17.2377 | -104.4627
17.2414 | -104.4508
17.2452 | -104.4391
17.2489 | -104.4272
17.2527 | -104.4154
17.2564 | -104.408
17.2602 | -104.4023
17.2639 | -104.3975
17.2677 | -104.3938
17.2714 | -104.3909
17.2752 | -104.3857
17.2789 | -104.3841
17.2827 | -104.3839
17.2865 | -104.3799
17.2902 | -104.3788
17.294 | -104.3783
17.2977 | -104.377
17.3015 | -104.3777
17.3052 | -104.3797
17.309 | -104.3805
17.3127 | -104.386
17.3165 | -104.3884
17.3202 | -104.3903
17.324 | -104.3955
17.3277 | -104.4026
17.3315 | -104.4077
17.3352 | -104.4134
17.339 | -104.4184
17.3427 | -104.4226
17.3465 | -104.4273
17.3502 | -104.4327
17.354 | -104.4405
17.3577 | -104.4513
17.3615 | -104.4621
17.3652 | -104.471
17.369 | -104.4793
17.3727 | -104.4877
17.3765 | -104.4957
17.3802 | -104.5053
17.384 | -104.5139
17.3877 | -104.5209
17.3915 | -104.5283
17.3952 | -104.5355
17.399 | -104.5402
17.4027 | -104.5434
17.4065 | -104.5484
17.4102 | -104.5513
17.414 | -104.5532
17.4177 | -104.554
17.4215 | -104.5521
17.4252 | -104.5494
17.429 | -104.5475
17.4327 | -104.546
17.4365 | -104.5446
17.4402 | -104.5408
17.444 | -104.5364
17.4477 | -104.5315
17.4515 | -104.5257
17.4552 | -104.519
17.459 | -104.5132
17.4627 | -104.5091
17.4665 | -104.5022
17.4702 | -104.4932
17.474 | -104.4804
17.4777 | -104.4684
17.4815 | -104.4563
17.4852 | -104.4425
17.489 | -104.4277
17.4927 | -104.4149
17.4965 | -104.4026
17.5002 | -104.3914
17.504 | -104.3756
17.5077 | -104.3595
17.5115 | -104.3436
17.5152 | -104.3279
17.519 | -104.314
17.5227 | -104.3012
17.5265 | -104.2862
17.5302 | -104.2707
17.534 | -104.2594
17.5377 | -104.2475
17.5415 | -104.2356
17.5452 | -104.2229
17.549 | -104.2102
17.5527 | -104.1969
17.5565 | -104.1841
17.5602 | -104.1703
17.564 | -104.1555
17.5677 | -104.1408
17.5715 | -104.1246
17.5752 | -104.1068
17.579 | -104.0907
17.5827 | -104.0747
17.5865 | -104.0592
17.5902 | -104.0439
17.594 | -104.0274
17.5978 | -104.0122
17.6015 | -103.9984
17.6053 | -103.9827
17.609 | -103.9664
17.6128 | -103.9484
17.6165 | -103.9307
17.6203 | -103.9156
17.624 | -103.9022
17.6278 | -103.8875
17.6315 | -103.872
17.6353 | -103.8553
17.639 | -103.8402
17.6428 | -103.8263
17.6465 | -103.8135
17.6503 | -103.8036
17.654 | -103.7949
17.6578 | -103.7847
17.6615 | -103.7752
17.6653 | -103.7682
17.669 | -103.7618
17.6728 | -103.7562
17.6765 | -103.7494
17.6803 | -103.7405
17.684 | -103.7345
17.6878 | -103.7308
17.6915 | -103.7282
17.6953 | -103.7254
17.699 | -103.7235
17.7028 | -103.7241
17.7065 | -103.7262
17.7103 | -103.7276
17.714 | -103.7283
17.7178 | -103.731
17.7215 | -103.7345
17.7253 | -103.7381
17.729 | -103.7426
17.7328 | -103.7465
17.7365 | -103.7495
17.7403 | -103.7531
17.744 | -103.7531
17.7478 | -103.7539
17.7515 | -103.7562
17.7553 | -103.7595
17.759 | -103.7604
17.7628 | -103.7604
17.7665 | -103.7622
17.7703 | -103.7654
17.774 | -103.7657
17.7778 | -103.7647
17.7815 | -103.7639
17.7853 | -103.7621
17.789 | -103.7616
17.7928 | -103.7598
17.7965 | -103.7583
17.8003 | -103.756
17.804 | -103.7564
17.8078 | -103.7562
17.8115 | -103.7536
17.8153 | -103.7518
17.819 | -103.7513
17.8228 | -103.7527
17.8265 | -103.7514
17.8303 | -103.7497
17.834 | -103.7494
17.8378 | -103.7471
17.8415 | -103.7459
17.8453 | -103.7461
17.849 | -103.7428
17.8528 | -103.7475
17.8565 | -103.7462
17.8603 | -103.7478
17.864 | -103.7503
17.8678 | -103.7529
17.8715 | -103.7575
17.8753 | -103.7625
17.879 | -103.7662
17.8828 | -103.7722
17.8865 | -103.7804
17.8903 | -103.7881
17.894 | -103.7927
17.8978 | -103.7984
17.9015 | -103.8032
17.9053 | -103.8081
17.909 | -103.8145
17.9128 | -103.8207
17.9165 | -103.827
17.9203 | -103.8338
17.9241 | -103.8405
17.9278 | -103.8462
17.9316 | -103.8535
17.9353 | -103.8604
17.9391 | -103.8607
17.9428 | -103.8601
17.9466 | -103.8619
17.9503 | -103.8622
17.9541 | -103.8609
17.9578 | -103.857
17.9616 | -103.8515
17.9653 | -103.8461
17.9691 | -103.8427
17.9728 | -103.839
17.9766 | -103.8293
17.9803 | -103.8187
17.9841 | -103.8065
17.9878 | -103.7971
17.9916 | -103.7874
17.9953 | -103.7752
17.9991 | -103.7654
18.0028 | -103.7555
18.0066 | -103.7425
18.0103 | -103.7291
18.0141 | -103.7159
18.0178 | -103.7028
18.0216 | -103.6871
18.0253 | -103.6693
18.0291 | -103.6548
18.0328 | -103.6417
18.0366 | -103.6266
18.0403 | -103.6126
18.0441 | -103.5966
18.0478 | -103.5802
18.0516 | -103.5641
18.0553 | -103.5489
18.0591 | -103.5316
18.0628 | -103.5182
18.0666 | -103.5057
18.0703 | -103.4933
18.0741 | -103.4797
18.0778 | -103.4649
18.0816 | -103.4505
18.0853 | -103.4353
18.0891 | -103.4194
18.0928 | -103.408
18.0966 | -103.3963
18.1003 | -103.3845
18.1041 | -103.3766
18.1078 | -103.3684
18.1116 | -103.3586
18.1153 | -103.3475
18.1191 | -103.3378
18.1228 | -103.3307
18.1266 | -103.3221
18.1303 | -103.3129
18.1341 | -103.3071
18.1378 | -103.3038
18.1416 | -103.3006
18.1453 | -103.296
18.1491 | -103.2941
18.1528 | -103.2942
18.1566 | -103.2919
18.1603 | -103.2909
18.1641 | -103.2904
18.1678 | -103.2914
18.1716 | -103.2935
18.1753 | -103.2958
18.1791 | -103.2969
18.1828 | -103.2982
18.1866 | -103.2992
18.1903 | -103.3031
18.1941 | -103.3073
18.1978 | -103.3126
18.2016 | -103.3167
18.2053 | -103.3183
18.2091 | -103.3225
18.2128 | -103.3281
18.2166 | -103.3319
18.2203 | -103.3347
18.2241 | -103.3376
18.2278 | -103.339
18.2316 | -103.3412
18.2353 | -103.3447
18.2391 | -103.3479
18.2429 | -103.348
18.2466 | -103.3446
18.2504 | -103.3427
18.2541 | -103.3395
18.2579 | -103.3357
18.2616 | -103.3311
18.2654 | -103.3276
18.2691 | -103.325
18.2729 | -103.3227
18.2766 | -103.3181
18.2804 | -103.3112
18.2841 | -103.3075
18.2879 | -103.3039
18.2916 | -103.2987
18.2954 | -103.2942
18.2991 | -103.2876
18.3029 | -103.2803
18.3066 | -103.2725
18.3104 | -103.2664
18.3141 | -103.2608
18.3179 | -103.2543
18.3216 | -103.2425
18.3254 | -103.2316
18.3291 | -103.2197
18.3329 | -103.2111
18.3366 | -103.2037
18.3404 | -103.1926
18.3441 | -103.1788
18.3479 | -103.1683
18.3516 | -103.1556
18.3554 | -103.1399
18.3591 | -103.1258
18.3629 | -103.1087
18.3666 | -103.0915
18.3704 | -103.0764
18.3741 | -103.0619
18.3779 | -103.0441
18.3816 | -103.0236
18.3854 | -103.0052
18.3891 | -102.9875
18.3929 | -102.9635
18.3966 | -102.942
18.4004 | -102.9224
18.4041 | -102.9021
18.4079 | -102.8842
18.4116 | -102.8666
18.4154 | -102.849
18.4191 | -102.832
18.4229 | -102.8149
18.4266 | -102.8
18.4304 | -102.7855
18.4341 | -102.7726
18.4379 | -102.7594
18.4416 | -102.7466
18.4454 | -102.7343
18.4491 | -102.7231
18.4529 | -102.7132
18.4566 | -102.7019
18.4604 | -102.6945
18.4641 | -102.6871
18.4679 | -102.6789
18.4716 | -102.6722
18.4754 | -102.6662
18.4791 | -102.6632
18.4829 | -102.6601
18.4866 | -102.6589
18.4904 | -102.6595
18.4941 | -102.6605
18.4979 | -102.663
18.5016 | -102.6652
18.5054 | -102.6687
18.5091 | -102.675
18.5129 | -102.6798
18.5166 | -102.6853
18.5204 | -102.6912
18.5241 | -102.6978
18.5279 | -102.7041
18.5316 | -102.7119
18.5354 | -102.7206
18.5391 | -102.7304
18.5429 | -102.7413
18.5466 | -102.7509
18.5504 | -102.7597
18.5541 | -102.7688
18.5579 | -102.7797
18.5617 | -102.7916
18.5654 | -102.8033
18.5692 | -102.8148
18.5729 | -102.8248
18.5767 | -102.8351
18.5804 | -102.846
18.5842 | -102.8583
18.5879 | -102.8688
18.5917 | -102.8789
18.5954 | -102.8853
18.5992 | -102.893
18.6029 | -102.9048
18.6067 | -102.9154
18.6104 | -102.9233
18.6142 | -102.9305
18.6179 | -102.94
18.6217 | -102.9485
18.6254 | -102.9577
18.6292 | -102.9657
18.6329 | -102.9748
18.6367 | -102.9864
18.6404 | -102.9976
18.6442 | -103.0091
18.6479 | -103.0224
18.6517 | -103.0331
18.6554 | -103.0459
18.6592 | -103.0576
18.6629 | -103.0685
18.6667 | -103.0797
18.6704 | -103.0927
18.6742 | -103.1036
18.6779 | -103.115
18.6817 | -103.1288
18.6854 | -103.1437
18.6892 | -103.1579
18.6929 | -103.1713
18.6967 | -103.1862
18.7004 | -103.2002
18.7042 | -103.2138
18.7079 | -103.2299
18.7117 | -103.2465
18.7154 | -103.2637
18.7192 | -103.2779
18.7229 | -103.2936
18.7267 | -103.307
18.7304 | -103.3225
18.7342 | -103.3382
18.7379 | -103.3544
18.7417 | -103.3693
18.7454 | -103.3826
18.7492 | -103.3989
18.7529 | -103.413
18.7567 | -103.4242
18.7604 | -103.4353
18.7642 | -103.4485
18.7679 | -103.4608
18.7717 | -103.4698
18.7754 | -103.4779
18.7792 | -103.4864
18.7829 | -103.4951
18.7867 | -103.502
18.7904 | -103.5086
18.7942 | -103.5157
18.7979 | -103.5215
18.8017 | -103.5227
18.8054 | -103.5249
18.8092 | -103.5281
18.8129 | -103.5305
18.8167 | -103.5318
18.8204 | -103.5316
18.8242 | -103.5301
18.8279 | -103.5274
18.8317 | -103.5244
18.8354 | -103.521
18.8392 | -103.5165
18.8429 | -103.5092
18.8467 | -103.5032
18.8504 | -103.4985
18.8542 | -103.4941
18.8579 | -103.4889
18.8617 | -103.484
18.8654 | -103.4767
18.8692 | -103.4706
18.8729 | -103.4691
18.8767 | -103.4658
18.8804 | -103.462
18.8842 | -103.4595
18.888 | -103.4589
18.8917 | -103.4596
18.8955 | -103.462
18.8992 | -103.4634
18.903 | -103.466
18.9067 | -103.4698
18.9105 | -103.4755
18.9142 | -103.4801
18.918 | -103.4861
18.9217 | -103.495
18.9255 | -103.5051
18.9292 | -103.5153
18.933 | -103.5241
18.9367 | -103.5354
18.9405 | -103.5461
18.9442 | -103.5603
18.948 | -103.5741
18.9517 | -103.5867
18.9555 | -103.5993
18.9592 | -103.6124
18.963 | -103.6247
18.9667 | -103.6358
18.9705 | -103.6488
18.9742 | -103.662
18.978 | -103.6729
18.9817 | -103.6833
18.9855 | -103.6928
18.9892 | -103.7031
18.993 | -103.7146
18.9967 | -103.7245
19.0005 | -103.7317
19.0042 | -103.7395
19.008 | -103.7452
19.0117 | -103.7521
19.0155 | -103.7573
19.0192 | -103.7621
19.023 | -103.7673
19.0267 | -103.7737
19.0305 | -103.7751
19.0342 | -103.7782
19.038 | -103.7842
19.0417 | -103.791
19.0455 | -103.7922
19.0492 | -103.7948
19.053 | -103.7983
19.0567 | -103.8006
19.0605 | -103.8003
19.0642 | -103.7999
19.068 | -103.8001
19.0717 | -103.8017
19.0755 | -103.8004
19.0792 | -103.7976
19.083 | -103.7913
19.0867 | -103.7869
19.0905 | -103.7822
19.0942 | -103.7779
19.098 | -103.7733
19.1017 | -103.7673
19.1055 | -103.7613
19.1092 | -103.7549
19.113 | -103.749
19.1167 | -103.7407
19.1205 | -103.7285
19.1242 | -103.7158
19.128 | -103.7065
19.1317 | -103.6964
19.1355 | -103.6828
19.1392 | -103.6681
19.143 | -103.6504
19.1467 | -103.634
19.1505 | -103.6183
19.1542 | -103.6004
19.158 | -103.5822
19.1617 | -103.5644
19.1655 | -103.5452
19.1692 | -103.5259
19.173 | -103.5057
19.1767 | -103.4862
19.1805 | -103.4665
19.1842 | -103.4468
19.188 | -103.4257
19.1917 | -103.4008
19.1955 | -103.3771
19.1992 | -103.3557
19.203 | -103.3335
19.2068 | -103.3122
19.2105 | -103.2905
19.2143 | -103.2674
19.218 | -103.2436
19.2218 | -103.2214
19.2255 | -103.201
19.2293 | -103.1819
19.233 | -103.1645
19.2368 | -103.1481
19.2405 | -103.1311
19.2443 | -103.1142
19.248 | -103.0999
19.2518 | -103.0859
19.2555 | -103.0723
19.2593 | -103.0609
19.263 | -103.0498
19.2668 | -103.0383
19.2705 | -103.03
19.2743 | -103.0242
19.278 | -103.0169
19.2818 | -103.0112
19.2855 | -103.0084
19.2893 | -103.0064
19.293 | -103.0052
19.2968 | -103.0071
19.3005 | -103.0084
19.3043 | -103.0129
19.308 | -103.0158
19.3118 | -103.021
19.3155 | -103.0278
19.3193 | -103.037
19.323 | -103.0442
19.3268 | -103.0512
19.3305 | -103.0626
19.3343 | -103.0768
19.338 | -103.0913
19.3418 | -103.1062
19.3455 | -103.1202
19.3493 | -103.1355
19.353 | -103.1495
19.3568 | -103.1657
19.3605 | -103.1796
19.3643 | -103.1934
19.368 | -103.2057
19.3718 | -103.2152
19.3755 | -103.2241
19.3793 | -103.2324
19.383 | -103.2398
19.3868 | -103.2472
19.3905 | -103.2528
19.3943 | -103.2593
19.398 | -103.2656
19.4018 | -103.2687
19.4055 | -103.271
19.4093 | -103.274
19.413 | -103.2741
19.4168 | -103.2727
19.4205 | -103.269
19.4243 | -103.2669
19.428 | -103.2643
19.4318 | -103.2602
19.4355 | -103.2549
19.4393 | -103.2499
19.443 | -103.2449
19.4468 | -103.2391
19.4505 | -103.2325
19.4543 | -103.2259
19.458 | -103.2217
19.4618 | -103.22
19.4655 | -103.2181
19.4693 | -103.2134
19.473 | -103.2088
19.4768 | -103.2133
19.4805 | -103.212
19.4843 | -103.2118
19.488 | -103.2118
19.4918 | -103.2135
19.4955 | -103.2178
19.4993 | -103.2232
19.503 | -103.2267
19.5068 | -103.2296
19.5105 | -103.2356
19.5143 | -103.2427
19.518 | -103.2512
19.5218 | -103.2607
19.5256 | -103.269
19.5293 | -103.2793
19.5331 | -103.2918
19.5368 | -103.3032
19.5406 | -103.3144
19.5443 | -103.325
19.5481 | -103.3345
19.5518 | -103.3428
19.5556 | -103.3549
19.5593 | -103.3668
19.5631 | -103.3755
19.5668 | -103.3853
19.5706 | -103.3941
19.5743 | -103.4022
19.5781 | -103.4106
19.5818 | -103.4185
19.5856 | -103.4267
19.5893 | -103.4285
19.5931 | -103.4303
19.5968 | -103.4342
19.6006 | -103.4374
19.6043 | -103.4382
19.6081 | -103.4344
19.6118 | -103.4323
19.6156 | -103.431
19.6193 | -103.4286
19.6231 | -103.4294
19.6268 | -103.4304
19.6306 | -103.4322
19.6343 | -103.4308
19.6381 | -103.4301
19.6418 | -103.4298
19.6456 | -103.4275
19.6493 | -103.4257
19.6531 | -103.4243
19.6568 | -103.4224
19.6606 | -103.4214
19.6643 | -103.4192
19.6681 | -103.4178
19.6718 | -103.4151
19.6756 | -103.4128
19.6793 | -103.4108
19.6831 | -103.41
19.6868 | -103.4118
19.6906 | -103.4131
19.6943 | -103.4154
19.6981 | -103.4181
19.7018 | -103.4171
19.7056 | -103.4178
19.7093 | -103.4201
19.7131 | -103.4249
19.7168 | -103.43
19.7206 | -103.4342
19.7243 | -103.4399
19.7281 | -103.4465
19.7318 | -103.453
19.7356 | -103.4562
19.7393 | -103.4593
19.7431 | -103.4654
19.7468 | -103.4672
19.7506 | -103.4684
19.7543 | -103.4707
19.7581 | -103.4742
19.7618 | -103.4778
19.7656 | -103.4779
19.7693 | -103.4779
19.7731 | -103.478
19.7768 | -103.4758
19.7806 | -103.4726
19.7843 | -103.4713
19.7881 | -103.4706
19.7918 | -103.4669
19.7956 | -103.4634
19.7993 | -103.46
19.8031 | -103.4558
19.8068 | -103.4529
19.8106 | -103.4489
19.8143 | -103.4419
19.8181 | -103.4343
19.8218 | -103.4258
19.8256 | -103.4184
19.8293 | -103.4123
19.8331 | -103.4056
19.8368 | -103.395
19.8406 | -103.3856
19.8444 | -103.3789
19.8481 | -103.3725
19.8519 | -103.3657
19.8556 | -103.3567
19.8594 | -103.3519
19.8631 | -103.3471
19.8669 | -103.3424
19.8706 | -103.3406
19.8744 | -103.3363
19.8781 | -103.3291
19.8819 | -103.3217
19.8856 | -103.3163
19.8894 | -103.3139
19.8931 | -103.3079
19.8969 | -103.3
19.9006 | -103.2936
19.9044 | -103.2889
19.9081 | -103.2844
19.9119 | -103.2787
19.9156 | -103.2727
19.9194 | -103.2676
19.9231 | -103.2621
19.9269 | -103.2592
19.9306 | -103.2569
19.9344 | -103.2524
19.9381 | -103.2494
19.9419 | -103.2449
19.9456 | -103.2386
19.9494 | -103.2318
19.9531 | -103.2287
19.9569 | -103.2257
19.9606 | -103.2229
19.9644 | -103.2195
19.9681 | -103.215
19.9719 | -103.2133
19.9756 | -103.2137
19.9794 | -103.2147
19.9831 | -103.2174
19.9869 | -103.2208
19.9906 | -103.2237
19.9944 | -103.2257
19.9981 | -103.2284
20.0019 | -103.2325
20.0056 | -103.236
20.0094 | -103.2396
20.0131 | -103.2432
20.0169 | -103.2518
20.0206 | -103.2588
20.0244 | -103.266
20.0281 | -103.2731
20.0319 | -103.2835
20.0356 | -103.2927
20.0394 | -103.3008
20.0431 | -103.3067
20.0469 | -103.3121
20.0506 | -103.3196
20.0544 | -103.3258
20.0581 | -103.3293
20.0619 | -103.3334
20.0656 | -103.3356
20.0694 | -103.338
20.0731 | -103.3398
20.0769 | -103.3404
20.0806 | -103.3399
20.0844 | -103.338
20.0881 | -103.3338
20.0919 | -103.3288
20.0956 | -103.3236
20.0994 | -103.3198
20.1031 | -103.3151
20.1069 | -103.3095
20.1106 | -103.3017
20.1144 | -103.29
20.1181 | -103.2789
20.1219 | -103.2677
20.1256 | -103.2541
20.1294 | -103.2417
20.1331 | -103.2307
20.1369 | -103.2186
20.1406 | -103.2075
20.1444 | -103.1954
20.1481 | -103.1827
20.1519 | -103.1697
20.1556 | -103.1592
20.1594 | -103.1521
20.1632 | -103.1464
20.1669 | -103.1418
20.1707 | -103.1373
20.1744 | -103.1327
20.1782 | -103.1301
20.1819 | -103.1288
20.1857 | -103.1277
20.1894 | -103.1269
20.1932 | -103.1277
20.1969 | -103.1309
20.2007 | -103.1379
20.2044 | -103.1438
20.2082 | -103.1506
20.2119 | -103.1586
20.2157 | -103.1698
20.2194 | -103.1827
20.2232 | -103.1945
20.2269 | -103.2095
20.2307 | -103.2275
20.2344 | -103.2459
20.2382 | -103.2627
20.2419 | -103.2811
20.2457 | -103.2988
20.2494 | -103.3179
20.2532 | -103.3383
20.2569 | -103.3588
20.2607 | -103.3782
20.2644 | -103.3984
20.2682 | -103.4192
20.2719 | -103.4408
20.2757 | -103.4621
20.2794 | -103.4843
20.2832 | -103.5083
20.2869 | -103.5331
20.2907 | -103.5582
20.2944 | -103.5828
20.2982 | -103.6048
20.3019 | -103.6259
20.3057 | -103.6468
20.3094 | -103.6687
20.3132 | -103.6905
20.3169 | -103.7097
20.3207 | -103.7297
20.3244 | -103.7502
20.3282 | -103.7686
20.3319 | -103.7876
20.3357 | -103.8058
20.3394 | -103.8223
20.3432 | -103.8396
20.3469 | -103.8583
20.3507 | -103.8744
20.3544 | -103.8896
20.3582 | -103.9028
20.3619 | -103.9161
20.3657 | -103.929
20.3694 | -103.9409
20.3732 | -103.952
20.3769 | -103.9633
20.3807 | -103.9731
20.3844 | -103.9819
20.3882 | -103.9887
20.3919 | -103.9952
20.3957 | -104.0022
20.3994 | -104.008
20.4032 | -104.0125
20.4069 | -104.0141
20.4107 | -104.0158
20.4144 | -104.0176
20.4182 | -104.0144
20.4219 | -104.01
20.4257 | -104.0078
20.4294 | -104.0038
20.4332 | -103.9977
20.4369 | -103.9917
20.4407 | -103.9846
20.4444 | -103.9744
20.4482 | -103.9628
20.4519 | -103.9501
20.4557 | -103.9353
20.4594 | -103.9198
20.4632 | -103.9036
20.4669 | -103.8858
20.4707 | -103.8692
20.4744 | -103.8523
20.4782 | -103.8345
20.482 | -103.8155
20.4857 | -103.7946
20.4895 | -103.7732
20.4932 | -103.7537
20.497 | -103.7347
20.5007 | -103.7166
20.5045 | -103.6973
20.5082 | -103.6792
20.512 | -103.6606
20.5157 | -103.6423
20.5195 | -103.6247
20.5232 | -103.6071
20.527 | -103.5914
20.5307 | -103.5769
20.5345 | -103.5646
20.5382 | -103.5529
20.542 | -103.5425
20.5457 | -103.5342
20.5495 | -103.5271
20.5532 | -103.5199
20.557 | -103.5155
20.5607 | -103.5227
20.5645 | -103.5188
20.5682 | -103.5165
20.572 | -103.5152
20.5757 | -103.5139
20.5795 | -103.511
20.5832 | -103.5105
20.587 | -103.5126
20.5907 | -103.5172
20.5945 | -103.5224
20.5982 | -103.5258
20.602 | -103.528
20.6057 | -103.5309
20.6095 | -103.5369
20.6132 | -103.5421
20.617 | -103.5468
20.6207 | -103.5535
20.6245 | -103.5585
20.6282 | -103.5632
20.632 | -103.5688
20.6357 | -103.5743
20.6395 | -103.5783
20.6432 | -103.5824
20.647 | -103.5879
20.6507 | -103.5925
20.6545 | -103.5961
20.6582 | -103.5983
20.662 | -103.5992
20.6657 | -103.6014
20.6695 | -103.602
20.6732 | -103.6035
20.677 | -103.6021
20.6807 | -103.6004
20.6845 | -103.5974
20.6882 | -103.5916
20.692 | -103.5887
20.6957 | -103.5843
20.6995 | -103.579
20.7032 | -103.5722
20.707 | -103.5641
20.7107 | -103.5559
20.7145 | -103.5462
20.7182 | -103.5362
20.722 | -103.5255
20.7257 | -103.5132
20.7295 | -103.501
20.7332 | -103.4882
20.737 | -103.474
20.7407 | -103.4609
20.7445 | -103.4485
20.7482 | -103.4364
20.752 | -103.4243
20.7557 | -103.413
20.7595 | -103.4014
20.7632 | -103.3877
20.767 | -103.3733
20.7707 | -103.3587
20.7745 | -103.344
20.7782 | -103.3338
20.782 | -103.3238
20.7857 | -103.3139
20.7895 | -103.3037
20.7932 | -103.2935
20.797 | -103.2843
20.8008 | -103.2767
20.8045 | -103.2706
20.8083 | -103.2656
20.812 | -103.2617
20.8158 | -103.2597
20.8195 | -103.257
20.8233 | -103.2551
20.827 | -103.2515
20.8308 | -103.2487
20.8345 | -103.2486
20.8383 | -103.25
20.842 | -103.2507
20.8458 | -103.2521
20.8495 | -103.2546
20.8533 | -103.2553
20.857 | -103.2556
20.8608 | -103.2566
20.8645 | -103.2598
20.8683 | -103.2635
20.872 | -103.267
20.8758 | -103.2735
20.8795 | -103.2777
20.8833 | -103.2795
20.887 | -103.2825
20.8908 | -103.2852
20.8945 | -103.2876
20.8983 | -103.2898
20.902 | -103.2913
20.9058 | -103.291
20.9095 | -103.2926
20.9133 | -103.2932
20.917 | -103.2932
20.9208 | -103.2957
20.9245 | -103.2972
20.9283 | -103.2967
20.932 | -103.2956
20.9358 | -103.2963
20.9395 | -103.2958
20.9433 | -103.2941
20.947 | -103.2933
20.9508 | -103.2899
20.9545 | -103.2879
20.9583 | -103.286
20.962 | -103.281
20.9658 | -103.2784
20.9695 | -103.2733
20.9733 | -103.2693
20.977 | -103.2644
20.9808 | -103.2605
20.9845 | -103.2567
20.9883 | -103.2453
20.992 | -103.2335
20.9958 | -103.2253
20.9995 | -103.2147
21.0033 | -103.2023
21.007 | -103.1911
21.0108 | -103.18
21.0145 | -103.1669
21.0183 | -103.155
21.022 | -103.1448
21.0258 | -103.1321
21.0295 | -103.1199
21.0333 | -103.1092
21.037 | -103.0971
21.0408 | -103.0858
21.0445 | -103.0757
21.0483 | -103.0635
21.052 | -103.0498
21.0558 | -103.0376
21.0595 | -103.0285
21.0633 | -103.0203
21.067 | -103.012
21.0708 | -103.0035
21.0745 | -102.9922
21.0783 | -102.9808
21.082 | -102.9693
21.0858 | -102.9582
21.0895 | -102.9476
21.0933 | -102.9399
21.097 | -102.9332
21.1008 | -102.9362
21.1045 | -102.9312
21.1083 | -102.9243
21.112 | -102.9201
21.1158 | -102.9182
21.1196 | -102.9167
21.1233 | -102.9132
21.1271 | -102.9142
21.1308 | -102.915
21.1346 | -102.9136
21.1383 | -102.9136
21.1421 | -102.9145
21.1458 | -102.9157
21.1496 | -102.9173
21.1533 | -102.9198
21.1571 | -102.9225
21.1608 | -102.9269
21.1646 | -102.9301
21.1683 | -102.9351
21.1721 | -102.9393
21.1758 | -102.9424
21.1796 | -102.9468
21.1833 | -102.9507
21.1871 | -102.9524
21.1908 | -102.9526
21.1946 | -102.9541
21.1983 | -102.9562
21.2021 | -102.9565
21.2058 | -102.957
21.2096 | -102.9541
21.2133 | -102.9518
21.2171 | -102.9468
21.2208 | -102.9417
21.2246 | -102.9373
21.2283 | -102.9317
21.2321 | -102.9245
21.2358 | -102.9144
21.2396 | -102.9031
21.2433 | -102.8928
21.2471 | -102.883
21.2508 | -102.873
21.2546 | -102.8631
21.2583 | -102.8522
21.2621 | -102.8411
21.2658 | -102.8297
21.2696 | -102.8196
21.2733 | -102.8083
21.2771 | -102.7973
21.2808 | -102.7881
21.2846 | -102.7811
21.2883 | -102.7783
21.2921 | -102.774
21.2958 | -102.7666
21.2996 | -102.7617
21.3033 | -102.759
21.3071 | -102.7563
21.3108 | -102.7554
21.3146 | -102.7566
21.3183 | -102.7579
21.3221 | -102.7613
21.3258 | -102.7671
21.3296 | -102.7746
21.3333 | -102.7833
21.3371 | -102.7921
21.3408 | -102.801
21.3446 | -102.8093
21.3483 | -102.8204
21.3521 | -102.8338
21.3558 | -102.8492
21.3596 | -102.8638
21.3633 | -102.8803
21.3671 | -102.8967
21.3708 | -102.913
21.3746 | -102.9305
21.3783 | -102.9485
21.3821 | -102.9668
21.3858 | -102.9857
21.3896 | -103.0044
21.3933 | -103.0195
21.3971 | -103.0317
21.4008 | -103.0437
21.4046 | -103.0564
21.4083 | -103.0709
21.4121 | -103.0853
21.4158 | -103.0974
21.4196 | -103.1079
21.4233 | -103.1178
21.4271 | -103.1265
21.4308 | -103.1327
21.4346 | -103.1396
21.4383 | -103.1453
21.4421 | -103.1479
21.4459 | -103.1527
21.4496 | -103.1564
21.4534 | -103.1587
21.4571 | -103.1635
21.4609 | -103.167
21.4646 | -103.1676
21.4684 | -103.168
21.4721 | -103.1657
21.4759 | -103.165
21.4796 | -103.1627
21.4834 | -103.1621
21.4871 | -103.1614
21.4909 | -103.1587
21.4946 | -103.156
21.4984 | -103.1547
21.5021 | -103.1501
21.5059 | -103.1459
21.5096 | -103.1411
21.5134 | -103.1378
21.5171 | -103.1309
21.5209 | -103.1235
21.5246 | -103.1164
21.5284 | -103.1077
21.5321 | -103.0981
21.5359 | -103.0904
21.5396 | -103.085
21.5434 | -103.0797
21.5471 | -103.074
21.5509 | -103.0679
21.5546 | -103.0616
21.5584 | -103.0547
21.5621 | -103.0505
21.5659 | -103.0468
21.5696 | -103.0412
21.5734 | -103.036
21.5771 | -103.0321
21.5809 | -103.03
21.5846 | -103.027
21.5884 | -103.0237
21.5921 | -103.022
21.5959 | -103.0255
21.5996 | -103.0293
21.6034 | -103.0315
21.6071 | -103.0311
21.6109 | -103.0325
21.6146 | -103.0341
21.6184 | -103.0356
21.6221 | -103.0393
21.6259 | -103.0413
21.6296 | -103.0444
21.6334 | -103.0471
21.6371 | -103.0503
21.6409 | -103.0582
21.6446 | -103.0632
21.6484 | -103.0686
21.6521 | -103.0761
21.6559 | -103.0839
21.6596 | -103.0919
21.6634 | -103.0995
21.6671 | -103.1065
21.6709 | -103.1126
21.6746 | -103.1201
21.6784 | -103.1297
21.6821 | -103.1385
21.6859 | -103.1474
21.6896 | -103.1569
21.6934 | -103.1672
21.6971 | -103.1777
21.7009 | -103.1893
21.7046 | -103.2036
21.7084 | -103.2177
21.7121 | -103.231
21.7159 | -103.2441
21.7196 | -103.2591
21.7234 | -103.2738
21.7271 | -103.2878
21.7309 | -103.301
21.7346 | -103.3156
21.7384 | -103.3315
21.7421 | -103.3466
21.7459 | -103.3613
21.7496 | -103.375
21.7534 | -103.3877
21.7571 | -103.4001
21.7609 | -103.4095
21.7647 | -103.4166
21.7684 | -103.425
21.7722 | -103.4323
21.7759 | -103.4389
21.7797 | -103.4437
21.7834 | -103.4462
21.7872 | -103.4492
21.7909 | -103.4494
21.7947 | -103.4515
21.7984 | -103.4529
21.8022 | -103.4544
21.8059 | -103.4523
21.8097 | -103.4486
21.8134 | -103.4465
21.8172 | -103.443
21.8209 | -103.4377
21.8247 | -103.4354
21.8284 | -103.4339
21.8322 | -103.4272
21.8359 | -103.4205
21.8397 | -103.4151
21.8434 | -103.4065
21.8472 | -103.399
21.8509 | -103.3931
21.8547 | -103.3864
21.8584 | -103.3787
21.8622 | -103.3716
21.8659 | -103.3639
21.8697 | -103.3562
21.8734 | -103.3483
21.8772 | -103.3408
21.8809 | -103.3341
21.8847 | -103.3282
21.8884 | -103.3225
21.8922 | -103.3175
21.8959 | -103.3104
21.8997 | -103.3028
21.9034 | -103.2925
21.9072 | -103.282
21.9109 | -103.2739
21.9147 | -103.2667
21.9184 | -103.2584
21.9222 | -103.2494
21.9259 | -103.2416
21.9297 | -103.2342
21.9334 | -103.2289
21.9372 | -103.2218
21.9409 | -103.2121
21.9447 | -103.2044
21.9484 | -103.1975
21.9522 | -103.1885
21.9559 | -103.1792
21.9597 | -103.1687
21.9634 | -103.1562
21.9672 | -103.1448
21.9709 | -103.134
21.9747 | -103.1214
21.9784 | -103.1094
21.9822 | -103.0979
21.9859 | -103.0855
21.9897 | -103.0728
21.9934 | -103.0606
21.9972 | -103.0481
22.0009 | -103.033
22.0047 | -103.0175
22.0084 | -103.0008
22.0122 | -102.9835
22.0159 | -102.9659
22.0197 | -102.9487
22.0234 | -102.9334
22.0272 | -102.9199
22.0309 | -102.9045
22.0347 | -102.8886
22.0384 | -102.8742
22.0422 | -102.8624
22.0459 | -102.8503
22.0497 | -102.8362
22.0534 | -102.8233
22.0572 | -102.8127
22.0609 | -102.8028
22.0647 | -102.7925
22.0684 | -102.7826
22.0722 | -102.7765
22.0759 | -102.7719
22.0797 | -102.7669
22.0835 | -102.7639
22.0872 | -102.7596
22.091 | -102.7573
22.0947 | -102.7563
22.0985 | -102.7551
22.1022 | -102.7542
22.106 | -102.7529
22.1097 | -102.7535
22.1135 | -102.7543
22.1172 | -102.7561
22.121 | -102.7564
22.1247 | -102.7565
22.1285 | -102.7583
22.1322 | -102.7608
22.136 | -102.7614
22.1397 | -102.7622
22.1435 | -102.7662
22.1472 | -102.7697
22.151 | -102.771
22.1547 | -102.7738
22.1585 | -102.776
22.1622 | -102.7773
22.166 | -102.7786
22.1697 | -102.7803
22.1735 | -102.7841
22.1772 | -102.7877
22.181 | -102.7921
22.1847 | -102.7976
22.1885 | -102.8034
22.1922 | -102.8089
22.196 | -102.8134
22.1997 | -102.8192
22.2035 | -102.8243
22.2072 | -102.8274
22.211 | -102.8342
22.2147 | -102.8396
22.2185 | -102.8465
22.2222 | -102.8535
22.226 | -102.8597
22.2297 | -102.8663
22.2335 | -102.8713
22.2372 | -102.8769
22.241 | -102.8812
22.2447 | -102.8864
22.2485 | -102.8917
22.2522 | -102.898
22.256 | -102.9061
22.2597 | -102.9122
22.2635 | -102.9188
22.2672 | -102.9265
22.271 | -102.9346
22.2747 | -102.9437
22.2785 | -102.9531
22.2822 | -102.9613
22.286 | -102.9679
22.2897 | -102.9775
22.2935 | -102.9883
22.2972 | -102.9999
22.301 | -103.0119
22.3047 | -103.022
22.3085 | -103.0333
22.3122 | -103.044
22.316 | -103.0559
22.3197 | -103.0673
22.3235 | -103.0801
22.3272 | -103.094
22.331 | -103.1087
22.3347 | -103.122
22.3385 | -103.1329
22.3422 | -103.1433
22.346 | -103.1538
22.3497 | -103.1648
22.3535 | -103.1753
22.3572 | -103.1834
22.361 | -103.1893
22.3647 | -103.1966
22.3685 | -103.2017
22.3722 | -103.2062
22.376 | -103.2114
22.3797 | -103.2158
22.3835 | -103.2179
22.3872 | -103.216
22.391 | -103.2135
22.3947 | -103.2108
22.3985 | -103.2059
22.4022 | -103.2023
22.406 | -103.1982
22.4098 | -103.1913
22.4135 | -103.1834
22.4173 | -103.1747
22.421 | -103.1635
22.4248 | -103.1512
22.4285 | -103.1372
22.4323 | -103.1222
22.436 | -103.1097
22.4398 | -103.0955
22.4435 | -103.0798
22.4473 | -103.0627
22.451 | -103.046
22.4548 | -103.0292
22.4585 | -103.0112
22.4623 | -102.9931
22.466 | -102.9755
22.4698 | -102.9566
22.4735 | -102.9375
22.4773 | -102.921
22.481 | -102.905
22.4848 | -102.8882
22.4885 | -102.8727
22.4923 | -102.86
22.496 | -102.8463
22.4998 | -102.833
22.5035 | -102.8226
22.5073 | -102.8136
22.511 | -102.8062
22.5148 | -102.8007
22.5185 | -102.7978
22.5223 | -102.7977
22.526 | -102.8009
22.5298 | -102.8031
22.5335 | -102.8079
22.5373 | -102.8143
22.541 | -102.8235
22.5448 | -102.8346
22.5485 | -102.8479
22.5523 | -102.8644
22.556 | -102.8815
22.5598 | -102.9007
22.5635 | -102.9207
22.5673 | -102.9428
22.571 | -102.9675
22.5748 | -102.9935
22.5785 | -103.0208
22.5823 | -103.0486
22.586 | -103.0771
22.5898 | -103.1073
22.5935 | -103.1374
22.5973 | -103.1699
22.601 | -103.2036
22.6048 | -103.2369
22.6085 | -103.2689
22.6123 | -103.2997
22.616 | -103.3286
22.6198 | -103.3575
22.6235 | -103.3883
22.6273 | -103.4171
22.631 | -103.4478
22.6348 | -103.4765
22.6385 | -103.5057
22.6423 | -103.5326
22.646 | -103.5588
22.6498 | -103.5832
22.6535 | -103.6068
22.6573 | -103.6289
22.661 | -103.6495
22.6648 | -103.6696
22.6685 | -103.6893
22.6723 | -103.707
22.676 | -103.7237
22.6798 | -103.7408
22.6835 | -103.7543
22.6873 | -103.7662
22.691 | -103.7758
22.6948 | -103.7823
22.6985 | -103.7897
22.7023 | -103.7987
22.706 | -103.8055
22.7098 | -103.809
22.7135 | -103.8122
22.7173 | -103.8165
22.7211 | -103.8162
22.7248 | -103.8111
22.7286 | -103.8081
22.7323 | -103.8079
22.7361 | -103.8068
22.7398 | -103.8045
22.7436 | -103.8041
22.7473 | -103.8036
22.7511 | -103.8026
22.7548 | -103.8009
22.7586 | -103.798
22.7623 | -103.7949
22.7661 | -103.7915
22.7698 | -103.789
22.7736 | -103.7862
22.7773 | -103.7825
22.7811 | -103.7796
22.7848 | -103.7784
22.7886 | -103.7761
22.7923 | -103.772
22.7961 | -103.7673
22.7998 | -103.7625
22.8036 | -103.7601
22.8073 | -103.76
22.8111 | -103.7577
22.8148 | -103.7543
22.8186 | -103.7522
22.8223 | -103.7529
22.8261 | -103.7528
22.8298 | -103.753
22.8336 | -103.754
22.8373 | -103.752
22.8411 | -103.7492
22.8448 | -103.7449
22.8486 | -103.7436
22.8523 | -103.7431
22.8561 | -103.7424
22.8598 | -103.7417
22.8636 | -103.741
22.8673 | -103.7378
22.8711 | -103.7332
22.8748 | -103.7273
22.8786 | -103.7235
22.8823 | -103.7198
22.8861 | -103.7154
22.8898 | -103.709
22.8936 | -103.7012
22.8973 | -103.6943
22.9011 | -103.6887
22.9048 | -103.6822
22.9086 | -103.6753
22.9123 | -103.6699
22.9161 | -103.666
22.9198 | -103.6597
22.9236 | -103.6521
22.9273 | -103.6461
22.9311 | -103.64
22.9348 | -103.6325
22.9386 | -103.6259
22.9423 | -103.6208
22.9461 | -103.6138
22.9498 | -103.6076
22.9536 | -103.5983
22.9573 | -103.5911
22.9611 | -103.5822
22.9648 | -103.5728
22.9686 | -103.5611
22.9723 | -103.5495
22.9761 | -103.5389
22.9798 | -103.5291
22.9836 | -103.5197
22.9873 | -103.5115
22.9911 | -103.5027
22.9948 | -103.4934
22.9986 | -103.4845
23.0023 | -103.4752
23.0061 | -103.4638
23.0098 | -103.4519
23.0136 | -103.4404
23.0173 | -103.4304
23.0211 | -103.4189
23.0248 | -103.4088
23.0286 | -103.3983
23.0323 | -103.3883
23.0361 | -103.3763
23.0399 | -103.365
23.0436 | -103.3554
23.0474 | -103.3451
23.0511 | -103.3347
23.0549 | -103.3264
23.0586 | -103.3151
23.0624 | -103.3044
23.0661 | -103.293
23.0699 | -103.2816
23.0736 | -103.2706
23.0774 | -103.2589
23.0811 | -103.248
23.0849 | -103.2392
23.0886 | -103.2293
23.0924 | -103.219
23.0961 | -103.2087
23.0999 | -103.1968
23.1036 | -103.1856
23.1074 | -103.1768
23.1111 | -103.165
23.1149 | -103.1551
23.1186 | -103.1464
23.1224 | -103.1358
23.1261 | -103.1232
23.1299 | -103.1103
23.1336 | -103.0969
23.1374 | -103.0828
23.1411 | -103.0681
23.1449 | -103.0543
23.1486 | -103.0388
23.1524 | -103.0249
23.1561 | -103.0097
23.1599 | -102.9934
23.1636 | -102.9796
23.1674 | -102.9642
23.1711 | -102.9487
23.1749 | -102.932
23.1786 | -102.915
23.1824 | -102.8998
23.1861 | -102.8835
23.1899 | -102.8647
23.1936 | -102.8484
23.1974 | -102.8317
23.2011 | -102.814
23.2049 | -102.7947
23.2086 | -102.7746
23.2124 | -102.7552
23.2161 | -102.7352
23.2199 | -102.7157
23.2236 | -102.6964
23.2274 | -102.6751
23.2311 | -102.6546
23.2349 | -102.6363
23.2386 | -102.6158
23.2424 | -102.5939
23.2461 | -102.5745
23.2499 | -102.5556
23.2536 | -102.5363
23.2574 | -102.5167
23.2611 | -102.4999
23.2649 | -102.4845
23.2686 | -102.4684
23.2724 | -102.45
23.2761 | -102.4322
23.2799 | -102.416
23.2836 | -102.4013
23.2874 | -102.3866
23.2911 | -102.3734
23.2949 | -102.3618
23.2986 | -102.3513
23.3024 | -102.3414
23.3061 | -102.3307
23.3099 | -102.32
23.3136 | -102.3095
23.3174 | -102.2985
23.3211 | -102.2894
23.3249 | -102.2802
23.3286 | -102.2732
23.3324 | -102.2673
23.3361 | -102.2641
23.3399 | -102.2631
23.3436 | -102.2618
23.3474 | -102.2607
23.3511 | -102.261
23.3549 | -102.2617
23.3587 | -102.2636
23.3624 | -102.2662
23.3662 | -102.2701
23.3699 | -102.274
23.3737 | -102.2788
23.3774 | -102.2863
23.3812 | -102.2922
23.3849 | -102.2987
23.3887 | -102.3063
23.3924 | -102.3146
23.3962 | -102.3254
23.3999 | -102.3379
23.4037 | -102.3508
23.4074 | -102.3632
23.4112 | -102.3764
23.4149 | -102.3892
23.4187 | -102.4023
23.4224 | -102.4147
23.4262 | -102.4283
23.4299 | -102.444
23.4337 | -102.4599
23.4374 | -102.4742
23.4412 | -102.4908
23.4449 | -102.5068
23.4487 | -102.5228
23.4524 | -102.5404
23.4562 | -102.559
23.4599 | -102.5758
23.4637 | -102.593
23.4674 | -102.6126
23.4712 | -102.63
23.4749 | -102.6455
23.4787 | -102.6614
23.4824 | -102.6746
23.4862 | -102.6899
23.4899 | -102.7055
23.4937 | -102.719
23.4974 | -102.7326
23.5012 | -102.7463
23.5049 | -102.7602
23.5087 | -102.7738
23.5124 | -102.7854
23.5162 | -102.7953
23.5199 | -102.8068
23.5237 | -102.8174
23.5274 | -102.8267
23.5312 | -102.8361
23.5349 | -102.8458
23.5387 | -102.8548
23.5424 | -102.8627
23.5462 | -102.8718
23.5499 | -102.881
23.5537 | -102.8887
23.5574 | -102.8959
23.5612 | -102.9027
23.5649 | -102.9081
23.5687 | -102.9142
23.5724 | -102.921
23.5762 | -102.9261
23.5799 | -102.9281
23.5837 | -102.9301
23.5874 | -102.9356
23.5912 | -102.9405
23.5949 | -102.9459
23.5987 | -102.9502
23.6024 | -102.9533
23.6062 | -102.959
23.6099 | -102.9649
23.6137 | -102.9681
23.6174 | -102.9722
23.6212 | -102.9742
23.6249 | -102.9757
23.6287 | -102.9794
23.6324 | -102.9826
23.6362 | -102.9846
23.6399 | -102.9877
23.6437 | -102.9913
23.6474 | -102.9954
23.6512 | -103.0004
23.6549 | -103.003
23.6587 | -103.0064
23.6624 | -103.0082
23.6662 | -103.0094
23.6699 | -103.0112
23.6737 | -103.0127
23.6775 | -103.0136
23.6812 | -103.0148
23.685 | -103.0169
23.6887 | -103.017
23.6925 | -103.0197
23.6962 | -103.0215
23.7 | -103.0237
23.7037 | -103.0257
23.7075 | -103.0268
23.7112 | -103.0279
23.715 | -103.0306
23.7187 | -103.0331
23.7225 | -103.0352
23.7262 | -103.0373
23.73 | -103.0382
23.7337 | -103.036
23.7375 | -103.036
23.7412 | -103.0398
23.745 | -103.0416
23.7487 | -103.0412
23.7525 | -103.0428
23.7562 | -103.0456
23.76 | -103.0477
23.7637 | -103.0492
23.7675 | -103.0507
23.7712 | -103.0499
23.775 | -103.0505
23.7787 | -103.0518
23.7825 | -103.0536
23.7862 | -103.0551
23.79 | -103.0561
23.7937 | -103.0575
23.7975 | -103.0594
23.8012 | -103.0611
23.805 | -103.0667
23.8087 | -103.0705
23.8125 | -103.0753
23.8162 | -103.0798
23.82 | -103.0856
23.8237 | -103.0894
23.8275 | -103.0943
23.8312 | -103.1009
23.835 | -103.1067
23.8387 | -103.1114
23.8425 | -103.1151
23.8462 | -103.1207
23.85 | -103.1254
23.8537 | -103.1306
23.8575 | -103.136
23.8612 | -103.1419
23.865 | -103.1496
23.8687 | -103.1566
23.8725 | -103.1621
23.8762 | -103.169
23.88 | -103.1761
23.8837 | -103.1825
23.8875 | -103.1877
23.8912 | -103.1925
23.895 | -103.1982
23.8987 | -103.205
23.9025 | -103.2116
23.9062 | -103.2179
23.91 | -103.2241
23.9137 | -103.2299
23.9175 | -103.2383
23.9212 | -103.2459
23.925 | -103.25
23.9287 | -103.2527
23.9325 | -103.2565
23.9362 | -103.2602
23.94 | -103.265
23.9437 | -103.2694
23.9475 | -103.2723
23.9512 | -103.2765
23.955 | -103.2793
23.9587 | -103.2837
23.9625 | -103.2877
23.9662 | -103.2893
23.97 | -103.2916
23.9737 | -103.294
23.9775 | -103.2964
23.9812 | -103.2999
23.985 | -103.303
23.9887 | -103.3055
23.9925 | -103.3084
23.9962 | -103.3108
24 | -103.313
24.0038 | -103.3169
24.0075 | -103.3186
24.0113 | -103.32
24.015 | -103.3226
24.0188 | -103.3261
24.0225 | -103.3294
24.0263 | -103.3359
24.03 | -103.3403
24.0338 | -103.3461
24.0375 | -103.3508
24.0413 | -103.3552
24.045 | -103.3599
24.0488 | -103.3652
24.0525 | -103.3712
24.0563 | -103.3762
24.06 | -103.3843
24.0638 | -103.3922
24.0675 | -103.3992
24.0713 | -103.4081
24.075 | -103.4153
24.0788 | -103.4227
24.0825 | -103.431
24.0863 | -103.4414
24.09 | -103.4507
24.0938 | -103.4601
24.0975 | -103.4694
24.1013 | -103.4782
24.105 | -103.4863
24.1088 | -103.4973
24.1125 | -103.5068
24.1163 | -103.5174
24.12 | -103.5273
24.1238 | -103.5352
24.1275 | -103.5437
24.1313 | -103.5541
24.135 | -103.5643
24.1388 | -103.5747
24.1425 | -103.5839
24.1463 | -103.5942
24.15 | -103.6033
24.1538 | -103.6117
24.1575 | -103.6203
24.1613 | -103.6279
24.165 | -103.6344
24.1688 | -103.6425
24.1725 | -103.6495
24.1763 | -103.6573
24.18 | -103.6651
24.1838 | -103.672
24.1875 | -103.6756
24.1913 | -103.681
24.195 | -103.6891
24.1988 | -103.6957
24.2025 | -103.7006
24.2063 | -103.7033
24.21 | -103.7063
24.2138 | -103.7106
24.2175 | -103.7145
24.2213 | -103.718
24.225 | -103.7197
24.2288 | -103.7205
24.2325 | -103.722
24.2363 | -103.7232
24.24 | -103.7234
24.2438 | -103.7231
24.2475 | -103.7195
24.2513 | -103.7134
24.255 | -103.7016
24.2588 | -103.6854
24.2625 | -103.6727
24.2663 | -103.6558
24.27 | -103.6276
24.2738 | -103.5911
24.2775 | -103.5531
24.2813 | -103.5156
24.285 | -103.4835
24.2888 | -103.4526
24.2925 | -103.4184
24.2963 | -103.3843
24.3 | -103.3494
24.3038 | -103.3098
24.3075 | -103.2683
24.3113 | -103.225
24.315 | -103.1806
24.3188 | -103.135
24.3225 | -103.0856
24.3263 | -103.034
24.3301 | -102.9815
24.3338 | -102.9285
24.3376 | -102.8739
24.3413 | -102.8183
24.3451 | -102.7611
24.3488 | -102.6946
24.3526 | -102.6376
24.3563 | -102.581
24.3601 | -102.5251
24.3638 | -102.4692
24.3676 | -102.4141
24.3713 | -102.359
24.3751 | -102.3062
24.3788 | -102.2547
24.3826 | -102.2062
24.3863 | -102.1597
24.3901 | -102.1149
24.3938 | -102.0708
24.3976 | -102.0288
24.4013 | -101.9904
24.4051 | -101.9538
24.4088 | -101.9194
24.4126 | -101.8878
24.4163 | -101.8598
24.4201 | -101.8368
24.4238 | -101.8169
24.4276 | -101.8009
24.4313 | -101.7877
24.4351 | -101.7772
24.4388 | -101.7704
24.4426 | -101.7677
24.4463 | -101.7699
24.4501 | -101.7749
24.4538 | -101.7831
24.4576 | -101.7946
24.4613 | -101.8093
24.4651 | -101.8267
24.4688 | -101.8489
24.4726 | -101.8738
24.4763 | -101.9012
24.4801 | -101.9314
24.4838 | -101.9623
24.4876 | -101.9961
24.4913 | -102.0332
24.4951 | -102.0728
24.4988 | -102.1157
24.5026 | -102.1601
24.5063 | -102.2059
24.5101 | -102.2528
24.5138 | -102.2975
24.5176 | -102.3438
24.5213 | -102.3906
24.5251 | -102.4379
24.5288 | -102.4863
24.5326 | -102.536
24.5363 | -102.5844
24.5401 | -102.6294
24.5438 | -102.6755
24.5476 | -102.7193
24.5513 | -102.762
24.5551 | -102.804
24.5588 | -102.8416
24.5626 | -102.8731
24.5663 | -102.8902
24.5701 | -102.8948
24.5738 | -102.8813
24.5776 | -102.8307
24.5813 | -102.7835
24.5851 | -102.745
24.5888 | -102.7163
24.5926 | -102.6896
24.5963 | -102.6591
24.6001 | -102.6246
24.6038 | -102.5866
24.6076 | -102.5385
24.6113 | -102.4879
24.6151 | -102.439
24.6188 | -102.3866
24.6226 | -102.3261
24.6263 | -102.2562
24.6301 | -102.1761
24.6338 | -102.0916
24.6376 | -102.0036
24.6413 | -101.9102
24.6451 | -101.8125
24.6489 | -101.713
24.6526 | -101.6131
24.6564 | -101.5129
24.6601 | -101.4122
24.6639 | -101.3114
24.6676 | -101.2116
24.6714 | -101.1135
24.6751 | -101.0165
24.6789 | -100.9199
24.6826 | -100.8257
24.6864 | -100.7339
24.6901 | -100.6455
24.6939 | -100.5599
24.6976 | -100.478
24.7014 | -100.4002
24.7051 | -100.3262
24.7089 | -100.2559
24.7126 | -100.1903
24.7164 | -100.1296
24.7201 | -100.0737
24.7239 | -100.0225
24.7276 | -99.9773
24.7314 | -99.9374
24.7351 | -99.9034
24.7389 | -99.8758
24.7426 | -99.8534
24.7464 | -99.8376
24.7501 | -99.8272
24.7539 | -99.8232
24.7576 | -99.8262
24.7614 | -99.8346
24.7651 | -99.8479
24.7689 | -99.8661
24.7726 | -99.8902
24.7764 | -99.9199
24.7801 | -99.9553
24.7839 | -99.9978
24.7876 | -100.0458
24.7914 | -100.0971
24.7951 | -100.1526
24.7989 | -100.2141
24.8026 | -100.2797
24.8064 | -100.3489
24.8101 | -100.4213
24.8139 | -100.4982
24.8176 | -100.578
24.8214 | -100.6602
24.8251 | -100.7437
24.8289 | -100.8285
24.8326 | -100.9163
24.8364 | -101.0047
24.8401 | -101.0932
24.8439 | -101.1818
24.8476 | -101.2691
24.8514 | -101.3546
24.8551 | -101.4405
24.8589 | -101.5252
24.8626 | -101.6098
24.8664 | -101.6915
24.8701 | -101.7702
24.8739 | -101.8466
24.8776 | -101.9185
24.8814 | -101.9867
24.8851 | -102.0522
24.8889 | -102.0773
24.8926 | -102.1363
24.8964 | -102.19
24.9001 | -102.2418
24.9039 | -102.2912
24.9076 | -102.3365
24.9114 | -102.3806
24.9151 | -102.4234
24.9189 | -102.4719
24.9226 | -102.5247
24.9264 | -102.5836
24.9301 | -102.6613
24.9339 | -102.7187
24.9376 | -102.7532
24.9414 | -102.7686
24.9451 | -102.7775
24.9489 | -102.7852
24.9526 | -102.7931
24.9564 | -102.7986
24.9601 | -102.8052
24.9639 | -102.8097
24.9677 | -102.8142
24.9714 | -102.819
24.9752 | -102.8247
24.9789 | -102.8315
24.9827 | -102.8393
24.9864 | -102.8466
24.9902 | -102.8566
24.9939 | -102.8705
24.9977 | -102.8867
25.0014 | -102.9058
25.0052 | -102.9244
25.0089 | -102.9419
25.0127 | -102.9595
25.0164 | -102.978
25.0202 | -102.9979
25.0239 | -103.0164
25.0277 | -103.0358
25.0314 | -103.0571
25.0352 | -103.0797
25.0389 | -103.1037
25.0427 | -103.1278
25.0464 | -103.1512
25.0502 | -103.1755
25.0539 | -103.2001
25.0577 | -103.225
25.0614 | -103.2508
25.0652 | -103.2759
25.0689 | -103.3008
25.0727 | -103.3256
25.0764 | -103.3511
25.0802 | -103.376
25.0839 | -103.3994
25.0877 | -103.4211
25.0914 | -103.441
25.0952 | -103.4614
25.0989 | -103.4804
25.1027 | -103.4982
25.1064 | -103.5128
25.1102 | -103.528
25.1139 | -103.5449
25.1177 | -103.5614
25.1214 | -103.5753
25.1252 | -103.5877
25.1289 | -103.5993
25.1327 | -103.6089
25.1364 | -103.6149
25.1402 | -103.6202
25.1439 | -103.6264
25.1477 | -103.6301
25.1514 | -103.6314
25.1552 | -103.6334
25.1589 | -103.6352
25.1627 | -103.6361
25.1664 | -103.6356
25.1702 | -103.6342
25.1739 | -103.6315
25.1777 | -103.6293
25.1814 | -103.6269
25.1852 | -103.622
25.1889 | -103.6195
25.1927 | -103.6142
25.1964 | -103.6075
25.2002 | -103.6012
25.2039 | -103.5982
25.2077 | -103.5928
25.2114 | -103.5867
25.2152 | -103.5792
25.2189 | -103.5717
25.2227 | -103.5638
25.2264 | -103.554
25.2302 | -103.5472
25.2339 | -103.5428
25.2377 | -103.5372
25.2414 | -103.5319
25.2452 | -103.5286
25.2489 | -103.5253
25.2527 | -103.5225
25.2564 | -103.5202
25.2602 | -103.5178
25.2639 | -103.517
25.2677 | -103.5158
25.2714 | -103.5142
25.2752 | -103.5156
25.2789 | -103.5168
25.2827 | -103.5174
25.2865 | -103.5177
25.2902 | -103.5196
25.294 | -103.5228
25.2977 | -103.527
25.3015 | -103.5295
25.3052 | -103.5339
25.309 | -103.5391
25.3127 | -103.5454
25.3165 | -103.551
25.3202 | -103.5571
25.324 | -103.5623
25.3277 | -103.569
25.3315 | -103.5773
25.3352 | -103.5853
25.339 | -103.5919
25.3427 | -103.5993
25.3465 | -103.6061
25.3502 | -103.612
25.354 | -103.6185
25.3577 | -103.6242
25.3615 | -103.6292
25.3652 | -103.6345
25.369 | -103.6395
25.3727 | -103.646
25.3765 | -103.6512
25.3802 | -103.6562
25.384 | -103.6617
25.3877 | -103.6687
25.3915 | -103.6751
25.3952 | -103.6787
25.399 | -103.6812
25.4027 | -103.6856
25.4065 | -103.6887
25.4102 | -103.693
25.414 | -103.6972
25.4177 | -103.7008
25.4215 | -103.7047
25.4252 | -103.7076
25.429 | -103.7092
25.4327 | -103.7108
25.4365 | -103.7115
25.4402 | -103.7135
25.444 | -103.7172
25.4477 | -103.7202
25.4515 | -103.722
25.4552 | -103.7249
25.459 | -103.7259
25.4627 | -103.7283
25.4665 | -103.7336
25.4702 | -103.7349
25.474 | -103.7336
25.4777 | -103.7325
25.4815 | -103.732
25.4852 | -103.7322
25.489 | -103.7332
25.4927 | -103.7337
25.4965 | -103.7318
25.5002 | -103.7295
25.504 | -103.7293
25.5077 | -103.7308
25.5115 | -103.7318
25.5152 | -103.7315
25.519 | -103.7289
25.5227 | -103.7253
25.5265 | -103.7219
25.5302 | -103.7203
25.534 | -103.7199
25.5377 | -103.719
25.5415 | -103.7167
25.5452 | -103.7163
25.549 | -103.7136
25.5527 | -103.713
25.5565 | -103.714
25.5602 | -103.7128
25.564 | -103.7098
25.5677 | -103.7086
25.5715 | -103.7088
25.5752 | -103.7087
25.579 | -103.7062
25.5827 | -103.7045
25.5865 | -103.7044
25.5902 | -103.7028
25.594 | -103.6998
25.5978 | -103.698
25.6015 | -103.6962
25.6053 | -103.6935
25.609 | -103.6914
25.6128 | -103.69
25.6165 | -103.6883
25.6203 | -103.6859
25.624 | -103.6834
25.6278 | -103.6812
25.6315 | -103.6814
25.6353 | -103.6792
25.639 | -103.6767
25.6428 | -103.6748
25.6465 | -103.6743
25.6503 | -103.6733
25.654 | -103.673
25.6578 | -103.6708
25.6615 | -103.669
25.6653 | -103.6647
25.669 | -103.6618
25.6728 | -103.6587
25.6765 | -103.6563
25.6803 | -103.6543
25.684 | -103.6531
25.6878 | -103.6509
25.6915 | -103.6474
25.6953 | -103.6448
25.699 | -103.644
25.7028 | -103.6436
25.7065 | -103.6428
25.7103 | -103.6415
25.714 | -103.6398
25.7178 | -103.6379
25.7215 | -103.6369
25.7253 | -103.6378
25.729 | -103.6385
25.7328 | -103.6367
25.7365 | -103.6355
25.7403 | -103.6347
25.744 | -103.631
25.7478 | -103.6261
25.7515 | -103.6222
25.7553 | -103.6186
25.759 | -103.6151
25.7628 | -103.6113
25.7665 | -103.6088
25.7703 | -103.6035
25.774 | -103.5973
25.7778 | -103.591
25.7815 | -103.5848
25.7853 | -103.5786
25.789 | -103.5715
25.7928 | -103.5644
25.7965 | -103.5589
25.8003 | -103.5526
25.804 | -103.542
25.8078 | -103.5313
25.8115 | -103.5226
25.8153 | -103.5148
25.819 | -103.5059
25.8228 | -103.4981
25.8265 | -103.4878
25.8303 | -103.4762
25.834 | -103.4662
25.8378 | -103.4556
25.8415 | -103.445
25.8453 | -103.4349
25.849 | -103.4245
25.8528 | -103.4149
25.8565 | -103.4063
25.8603 | -103.3998
25.864 | -103.3925
25.8678 | -103.3829
25.8715 | -103.3727
25.8753 | -103.3621
25.879 | -103.3525
25.8828 | -103.3457
25.8865 | -103.3407
25.8903 | -103.3355
25.894 | -103.3274
25.8978 | -103.3198
25.9015 | -103.3133
25.9053 | -103.3086
25.909 | -103.3044
25.9128 | -103.3007
25.9165 | -103.2969
25.9203 | -103.2927
25.9241 | -103.29
25.9278 | -103.2888
25.9316 | -103.2871
25.9353 | -103.2853
25.9391 | -103.2851
25.9428 | -103.2878
25.9466 | -103.2901
25.9503 | -103.2923
25.9541 | -103.2947
25.9578 | -103.2988
25.9616 | -103.3027
25.9653 | -103.3068
25.9691 | -103.3172
25.9728 | -103.3241
25.9766 | -103.3313
25.9803 | -103.3383
25.9841 | -103.3459
25.9878 | -103.355
25.9916 | -103.3655
25.9953 | -103.3767
25.9991 | -103.3871
26.0028 | -103.3984
26.0066 | -103.4115
26.0103 | -103.4259
26.0141 | -103.4398
26.0178 | -103.4533
26.0216 | -103.4666
26.0253 | -103.4797
26.0291 | -103.4938
26.0328 | -103.51
26.0366 | -103.5261
26.0403 | -103.5429
26.0441 | -103.5564
26.0478 | -103.5702
26.0516 | -103.582
26.0553 | -103.5963
26.0591 | -103.6116
26.0628 | -103.6263
26.0666 | -103.6371
26.0703 | -103.6492
26.0741 | -103.6623
26.0778 | -103.6738
26.0816 | -103.6863
26.0853 | -103.6986
26.0891 | -103.7104
26.0928 | -103.7181
26.0966 | -103.7266
26.1003 | -103.7347
26.1041 | -103.7426
26.1078 | -103.7515
26.1116 | -103.7606
26.1153 | -103.7671
26.1191 | -103.7699
26.1228 | -103.7742
26.1266 | -103.7786
26.1303 | -103.7809
26.1341 | -103.7823
26.1378 | -103.7858
26.1416 | -103.7878
26.1453 | -103.7884
26.1491 | -103.7897
26.1528 | -103.7895
26.1566 | -103.7894
26.1603 | -103.7887
26.1641 | -103.7875
26.1678 | -103.7861
26.1716 | -103.7852
26.1753 | -103.782
26.1791 | -103.7785
26.1828 | -103.7747
26.1866 | -103.771
26.1903 | -103.7674
26.1941 | -103.7656
26.1978 | -103.7623
26.2016 | -103.7581
26.2053 | -103.7571
26.2091 | -103.7561
26.2128 | -103.755
26.2166 | -103.7524
26.2203 | -103.7484
26.2241 | -103.7459
26.2278 | -103.7414
26.2316 | -103.737
26.2353 | -103.7349
26.2391 | -103.7332
26.2429 | -103.7327
26.2466 | -103.7304
26.2504 | -103.7264
26.2541 | -103.7235
26.2579 | -103.7219
26.2616 | -103.722
26.2654 | -103.7225
26.2691 | -103.723
26.2729 | -103.7231
26.2766 | -103.7221
26.2804 | -103.7212
26.2841 | -103.722
26.2879 | -103.7235
26.2916 | -103.7246
26.2954 | -103.7244
26.2991 | -103.7251
26.3029 | -103.7283
26.3066 | -103.7303
26.3104 | -103.7328
26.3141 | -103.7363
26.3179 | -103.7381
26.3216 | -103.7406
26.3254 | -103.7426
26.3291 | -103.7459
26.3329 | -103.7492
26.3366 | -103.7521
26.3404 | -103.7537
26.3441 | -103.7554
26.3479 | -103.7584
26.3516 | -103.7618
26.3554 | -103.7642
26.3591 | -103.7658
26.3629 | -103.7668
26.3666 | -103.7693
26.3704 | -103.7712
26.3741 | -103.7718
26.3779 | -103.7711
26.3816 | -103.7711
26.3854 | -103.7726
26.3891 | -103.7745
26.3929 | -103.7772
26.3966 | -103.7793
26.4004 | -103.7791
26.4041 | -103.7722
26.4079 | -103.7661
26.4116 | -103.763
26.4154 | -103.7611
26.4191 | -103.7568
26.4229 | -103.7543
26.4266 | -103.7496
26.4304 | -103.7428
26.4341 | -103.7367
26.4379 | -103.7319
26.4416 | -103.7256
26.4454 | -103.7188
26.4491 | -103.7116
26.4529 | -103.7029
26.4566 | -103.6951
26.4604 | -103.6871
26.4641 | -103.6802
26.4679 | -103.67
26.4716 | -103.6596
26.4754 | -103.6485
26.4791 | -103.6374
26.4829 | -103.6256
26.4866 | -103.6149
26.4904 | -103.6046
26.4941 | -103.5919
26.4979 | -103.5792
26.5016 | -103.5644
26.5054 | -103.5498
26.5091 | -103.5344
26.5129 | -103.5184
26.5166 | -103.5037
26.5204 | -103.4901
26.5241 | -103.4763
26.5279 | -103.4624
26.5316 | -103.4481
26.5354 | -103.4341
26.5391 | -103.4206
26.5429 | -103.408
26.5466 | -103.3958
26.5504 | -103.3834
26.5541 | -103.3713
26.5579 | -103.3599
26.5617 | -103.349
26.5654 | -103.3383
26.5692 | -103.327
26.5729 | -103.3175
26.5767 | -103.3095
26.5804 | -103.3005
26.5842 | -103.2926
26.5879 | -103.2837
26.5917 | -103.2777
26.5954 | -103.2727
26.5992 | -103.2668
26.6029 | -103.2639
26.6067 | -103.2634
26.6104 | -103.2622
26.6142 | -103.2624
26.6179 | -103.262
26.6217 | -103.263
26.6254 | -103.2659
26.6292 | -103.2692
26.6329 | -103.2724
26.6367 | -103.2756
26.6404 | -103.2784
26.6442 | -103.2824
26.6479 | -103.287
26.6517 | -103.293
26.6554 | -103.2997
26.6592 | -103.307
26.6629 | -103.3152
26.6667 | -103.3232
26.6704 | -103.3317
26.6742 | -103.341
26.6779 | -103.3491
26.6817 | -103.3578
26.6854 | -103.3666
26.6892 | -103.3765
26.6929 | -103.3875
26.6967 | -103.3978
26.7004 | -103.4085
26.7042 | -103.4178
26.7079 | -103.4274
26.7117 | -103.4384
26.7154 | -103.4501
26.7192 | -103.4608
26.7229 | -103.4687
26.7267 | -103.4788
26.7304 | -103.4894
26.7342 | -103.5003
26.7379 | -103.5107
26.7417 | -103.5196
26.7454 | -103.5308
26.7492 | -103.543
26.7529 | -103.5563
26.7567 | -103.566
26.7604 | -103.5751
26.7642 | -103.5851
26.7679 | -103.5941
26.7717 | -103.6065
26.7754 | -103.6184
26.7792 | -103.6346
26.7829 | -103.6488
26.7867 | -103.6591
26.7904 | -103.6685
26.7942 | -103.6798
26.7979 | -103.6915
26.8017 | -103.7019
26.8054 | -103.7122
26.8092 | -103.7207
26.8129 | -103.7302
26.8167 | -103.7385
26.8204 | -103.746
26.8242 | -103.7549
26.8279 | -103.7613
26.8317 | -103.7676
26.8354 | -103.7745
26.8392 | -103.7816
26.8429 | -103.7897
26.8467 | -103.7966
26.8504 | -103.804
26.8542 | -103.8112
26.8579 | -103.819
26.8617 | -103.8254
26.8654 | -103.8322
26.8692 | -103.8389
26.8729 | -103.8443
26.8767 | -103.8527
26.8804 | -103.8589
26.8842 | -103.8633
26.888 | -103.8677
26.8917 | -103.8734
26.8955 | -103.8801
26.8992 | -103.8851
26.903 | -103.8893
26.9067 | -103.8944
26.9105 | -103.8999
26.9142 | -103.9048
26.918 | -103.9092
26.9217 | -103.9136
26.9255 | -103.9197
26.9292 | -103.9273
26.933 | -103.9317
26.9367 | -103.9348
26.9405 | -103.938
26.9442 | -103.9415
26.948 | -103.9461
26.9517 | -103.9518
26.9555 | -103.9577
26.9592 | -103.9639
26.963 | -103.9691
26.9667 | -103.9764
26.9705 | -103.9813
26.9742 | -103.9864
26.978 | -103.9916
26.9817 | -103.9954
26.9855 | -103.9991
26.9892 | -104.0046
26.993 | -104.0088
26.9967 | -104.0124
27.0005 | -104.017
27.0042 | -104.0228
27.008 | -104.0263
27.0117 | -104.0298
27.0155 | -104.0327
27.0192 | -104.0349
27.023 | -104.0375
27.0267 | -104.0409
27.0305 | -104.043
27.0342 | -104.0464
27.038 | -104.0485
27.0417 | -104.0491
27.0455 | -104.0523
27.0492 | -104.0555
27.053 | -104.0549
27.0567 | -104.0565
27.0605 | -104.0569
27.0642 | -104.058
27.068 | -104.0581
27.0717 | -104.0596
27.0755 | -104.0613
27.0792 | -104.0616
27.083 | -104.0614
27.0867 | -104.0623
27.0905 | -104.0626
27.0942 | -104.0617
27.098 | -104.0607
27.1017 | -104.0608
27.1055 | -104.0612
27.1092 | -104.0593
27.113 | -104.0586
27.1167 | -104.0584
27.1205 | -104.0573
27.1242 | -104.0545
27.128 | -104.0512
27.1317 | -104.0488
27.1355 | -104.0454
27.1392 | -104.0434
27.143 | -104.0409
27.1467 | -104.0381
27.1505 | -104.0342
27.1542 | -104.0263
27.158 | -104.0218
27.1617 | -104.0181
27.1655 | -104.0134
27.1692 | -104.0079
27.173 | -104.0034
27.1767 | -103.9962
27.1805 | -103.9869
27.1842 | -103.9792
27.188 | -103.9719
27.1917 | -103.9661
27.1955 | -103.9596
27.1992 | -103.9542
27.203 | -103.9482
27.2068 | -103.9398
27.2105 | -103.9335
27.2143 | -103.9257
27.218 | -103.9173
27.2218 | -103.9099
27.2255 | -103.9018
27.2293 | -103.8938
27.233 | -103.8862
27.2368 | -103.878
27.2405 | -103.8701
27.2443 | -103.8629
27.248 | -103.8564
27.2518 | -103.8499
27.2555 | -103.8418
27.2593 | -103.8346
27.263 | -103.8279
27.2668 | -103.8209
27.2705 | -103.8123
27.2743 | -103.8046
27.278 | -103.7951
27.2818 | -103.7878
27.2855 | -103.781
27.2893 | -103.774
27.293 | -103.7674
27.2968 | -103.7607
27.3005 | -103.7546
27.3043 | -103.7492
27.308 | -103.7432
27.3118 | -103.7379
27.3155 | -103.7345
27.3193 | -103.7322
27.323 | -103.7293
27.3268 | -103.7262
27.3305 | -103.7237
27.3343 | -103.7211
27.338 | -103.7184
27.3418 | -103.7153
27.3455 | -103.7145
27.3493 | -103.7132
27.353 | -103.7121
27.3568 | -103.7101
27.3605 | -103.7083
27.3643 | -103.7079
27.368 | -103.7069
27.3718 | -103.7065
27.3755 | -103.7069
27.3793 | -103.7092
27.383 | -103.7113
27.3868 | -103.7108
27.3905 | -103.7129
27.3943 | -103.7152
27.398 | -103.719
27.4018 | -103.7244
27.4055 | -103.7268
27.4093 | -103.7286
27.413 | -103.731
27.4168 | -103.7348
27.4205 | -103.737
27.4243 | -103.7406
27.428 | -103.7449
27.4318 | -103.7493
27.4355 | -103.7541
27.4393 | -103.7591
27.443 | -103.7633
27.4468 | -103.7674
27.4505 | -103.7718
27.4543 | -103.7778
27.458 | -103.781
27.4618 | -103.7844
27.4655 | -103.7884
27.4693 | -103.7934
27.473 | -103.7979
27.4768 | -103.8026
27.4805 | -103.808
27.4843 | -103.8117
27.488 | -103.8136
27.4918 | -103.8182
27.4955 | -103.8229
27.4993 | -103.8251
27.503 | -103.8274
27.5068 | -103.8312
27.5105 | -103.8357
27.5143 | -103.8385
27.518 | -103.8418
27.5218 | -103.8443
27.5256 | -103.847
27.5293 | -103.8489
27.5331 | -103.8507
27.5368 | -103.8527
27.5406 | -103.8572
27.5443 | -103.8604
27.5481 | -103.8628
27.5518 | -103.8653
27.5556 | -103.8682
27.5593 | -103.8713
27.5631 | -103.8758
27.5668 | -103.8816
27.5706 | -103.8854
27.5743 | -103.8884
27.5781 | -103.8919
27.5818 | -103.8961
27.5856 | -103.8979
27.5893 | -103.9006
27.5931 | -103.9084
27.5968 | -103.9147
27.6006 | -103.9203
27.6043 | -103.926
27.6081 | -103.9315
27.6118 | -103.9372
27.6156 | -103.9445
27.6193 | -103.9528
27.6231 | -103.9605
27.6268 | -103.9693
27.6306 | -103.9783
27.6343 | -103.9872
27.6381 | -103.9964
27.6418 | -104.0069
27.6456 | -104.0182
27.6493 | -104.0301
27.6531 | -104.0429
27.6568 | -104.057
27.6606 | -104.0729
27.6643 | -104.0879
27.6681 | -104.1031
27.6718 | -104.1173
27.6756 | -104.1312
27.6793 | -104.1451
27.6831 | -104.1601
27.6868 | -104.177
27.6906 | -104.1928
27.6943 | -104.209
27.6981 | -104.2254
27.7018 | -104.2413
27.7056 | -104.2579
27.7093 | -104.2748
27.7131 | -104.2919
27.7168 | -104.3092
27.7206 | -104.3257
27.7243 | -104.344
27.7281 | -104.3629
27.7318 | -104.3808
27.7356 | -104.3969
27.7393 | -104.4146
27.7431 | -104.4319
27.7468 | -104.451
27.7506 | -104.4689
27.7543 | -104.4846
27.7581 | -104.5005
27.7618 | -104.5156
27.7656 | -104.5301
27.7693 | -104.5428
27.7731 | -104.5556
27.7768 | -104.5691
27.7806 | -104.5791
27.7843 | -104.5887
27.7881 | -104.5966
27.7918 | -104.6026
27.7956 | -104.6109
27.7993 | -104.619
27.8031 | -104.6266
27.8068 | -104.6302
27.8106 | -104.6337
27.8143 | -104.6355
27.8181 | -104.6357
27.8218 | -104.6363
27.8256 | -104.6376
27.8293 | -104.6367
27.8331 | -104.6341
27.8368 | -104.6321
27.8406 | -104.6302
27.8444 | -104.625
27.8481 | -104.6181
27.8519 | -104.6127
27.8556 | -104.6057
27.8594 | -104.5954
27.8631 | -104.5859
27.8669 | -104.5758
27.8706 | -104.5629
27.8744 | -104.5515
27.8781 | -104.5413
27.8819 | -104.5294
27.8856 | -104.5172
27.8894 | -104.5052
27.8931 | -104.4928
27.8969 | -104.4795
27.9006 | -104.4675
27.9044 | -104.4539
27.9081 | -104.4382
27.9119 | -104.4244
27.9156 | -104.4119
27.9194 | -104.3995
27.9231 | -104.3835
27.9269 | -104.3629
27.9306 | -104.3407
27.9344 | -104.3222
27.9381 | -104.3058
27.9419 | -104.2902
27.9456 | -104.2759
27.9494 | -104.2633
27.9531 | -104.2502
27.9569 | -104.2367
27.9606 | -104.2255
27.9644 | -104.215
27.9681 | -104.2036
27.9719 | -104.1926
27.9756 | -104.1846
27.9794 | -104.1763
27.9831 | -104.1673
27.9869 | -104.1581
27.9906 | -104.1505
27.9944 | -104.1431
27.9981 | -104.1351
28.0019 | -104.1266