## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 272 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 275
Hood | 02-92B/P | 219
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1008.2 | hPa |  | PRC PRO 01
Humidity | 50.3 | %/Hr
Internal temperature of calibration bench | 24.36 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1017 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 25 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.15 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.97 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3122 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 20/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 24/07/2018
Technical Manager | Anthony HUE | Approval |  | 30/08/2018
CEO | Robert GLEMOT | Approval |  | 30/08/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4512 | 57.716
0.012209 | -36.0164 | 38.4786
0.018313 | -34.9484 | 28.2698
0.024417 | -34.4802 | 22.0543
0.030521 | -34.3548 | 18.2289
0.036626 | -34.2098 | 15.1681
0.04273 | -34.1126 | 12.8203
0.048834 | -34.0678 | 11.3152
0.054938 | -34.0048 | 10.1933
0.061043 | -33.9486 | 9.0998
0.067147 | -33.9364 | 8.2852
0.073251 | -33.9369 | 7.516
0.079355 | -33.9549 | 6.8116
0.08546 | -33.9668 | 6.1754
0.091564 | -33.9613 | 5.741
0.097668 | -33.9468 | 5.5164
0.10377 | -33.9329 | 5.2643
0.10988 | -33.9233 | 4.996
0.11598 | -33.9209 | 4.719
0.12209 | -33.9255 | 4.4495
0.12819 | -33.9323 | 4.1655
0.13429 | -33.9341 | 3.8763
0.1404 | -33.9328 | 3.6224
0.1465 | -33.9334 | 3.381
0.15261 | -33.9339 | 3.1679
0.15871 | -33.9331 | 2.9936
0.16482 | -33.9295 | 2.8666
0.17092 | -33.9268 | 2.8088
0.17702 | -33.9241 | 2.7266
0.18313 | -33.9208 | 2.6467
0.18923 | -33.9166 | 2.5399
0.19534 | -33.9108 | 2.4141
0.20144 | -33.9062 | 2.2679
0.20754 | -33.904 | 2.1387
0.21365 | -33.9025 | 2.0203
0.21975 | -33.9017 | 1.9225
0.22586 | -33.9021 | 1.8411
0.23196 | -33.9029 | 1.7817
0.23807 | -33.9044 | 1.7333
0.24417 | -33.9061 | 1.6803
0.25027 | -33.9084 | 1.6208
0.25638 | -33.9097 | 1.5573
0.26248 | -33.9113 | 1.4838
0.26859 | -33.9123 | 1.4074
0.27469 | -33.9129 | 1.3241
0.2808 | -33.9136 | 1.2444
0.2869 | -33.9138 | 1.1686
0.293 | -33.914 | 1.0992
0.29911 | -33.9139 | 1.0376
0.30521 | -33.9135 | 0.98063
0.31132 | -33.9134 | 0.94397
0.31742 | -33.9142 | 0.9012
0.32353 | -33.9146 | 0.86965
0.32963 | -33.9162 | 0.8431
0.33573 | -33.9185 | 0.82558
0.34184 | -33.9216 | 0.80541
0.34794 | -33.9261 | 0.783
0.35405 | -33.9302 | 0.7618
0.36015 | -33.9346 | 0.73778
0.36626 | -33.9382 | 0.7094
0.37236 | -33.9418 | 0.67118
0.37846 | -33.945 | 0.62246
0.38457 | -33.9454 | 0.55576
0.39067 | -33.9464 | 0.50202
0.39678 | -33.9469 | 0.4414
0.40288 | -33.9467 | 0.37811
0.40899 | -33.9458 | 0.31611
0.41509 | -33.9442 | 0.25726
0.42119 | -33.9434 | 0.21232
0.4273 | -33.9424 | 0.16124
0.4334 | -33.9415 | 0.11326
0.43951 | -33.9411 | 0.071651
0.44561 | -33.9416 | 0.033399
0.45172 | -33.9433 | 0.010889
0.45782 | -33.944 | -0.018263
0.46392 | -33.9453 | -0.04655
0.47003 | -33.9463 | -0.067384
0.47613 | -33.9471 | -0.09243
0.48224 | -33.9478 | -0.11549
0.48834 | -33.9476 | -0.1403
0.49445 | -33.9478 | -0.16484
0.50055 | -33.9477 | -0.18642
0.50665 | -33.9473 | -0.21104
0.51276 | -33.9465 | -0.23751
0.51886 | -33.9458 | -0.26508
0.52497 | -33.9447 | -0.28485
0.53107 | -33.9441 | -0.31004
0.53717 | -33.9436 | -0.33432
0.54328 | -33.9425 | -0.35544
0.54938 | -33.9423 | -0.37657
0.55549 | -33.9421 | -0.39753
0.56159 | -33.9424 | -0.41281
0.5677 | -33.9423 | -0.42827
0.5738 | -33.9428 | -0.44733
0.5799 | -33.9427 | -0.46426
0.58601 | -33.943 | -0.47744
0.59211 | -33.9438 | -0.48956
0.59822 | -33.9444 | -0.50389
0.60432 | -33.9449 | -0.52236
0.61043 | -33.9452 | -0.53684
0.61653 | -33.9457 | -0.55519
0.62263 | -33.9465 | -0.57074
0.62874 | -33.9466 | -0.58854
0.63484 | -33.9469 | -0.60473
0.64095 | -33.9469 | -0.62222
0.64705 | -33.9469 | -0.63842
0.65316 | -33.9469 | -0.6558
0.65926 | -33.947 | -0.67614
0.66536 | -33.9466 | -0.69552
0.67147 | -33.9466 | -0.71307
0.67757 | -33.9466 | -0.73119
0.68368 | -33.9464 | -0.75285
0.68978 | -33.9462 | -0.77106
0.69589 | -33.9459 | -0.79168
0.70199 | -33.9458 | -0.81246
0.70809 | -33.9455 | -0.8311
0.7142 | -33.9453 | -0.85097
0.7203 | -33.9449 | -0.86896
0.72641 | -33.945 | -0.89132
0.73251 | -33.9452 | -0.90981
0.73862 | -33.9458 | -0.93009
0.74472 | -33.946 | -0.94807
0.75082 | -33.9461 | -0.96972
0.75693 | -33.9468 | -0.99106
0.76303 | -33.9477 | -1.0129
0.76914 | -33.9498 | -1.033
0.77524 | -33.9511 | -1.0509
0.78135 | -33.9526 | -1.0697
0.78745 | -33.9543 | -1.0894
0.79355 | -33.9556 | -1.1092
0.79966 | -33.9576 | -1.129
0.80576 | -33.9602 | -1.1502
0.81187 | -33.9622 | -1.1723
0.81797 | -33.9646 | -1.1947
0.82408 | -33.9666 | -1.2171
0.83018 | -33.9686 | -1.2378
0.83628 | -33.9709 | -1.2586
0.84239 | -33.9726 | -1.2819
0.84849 | -33.9742 | -1.3049
0.8546 | -33.976 | -1.3268
0.8607 | -33.9775 | -1.3495
0.86681 | -33.9789 | -1.3688
0.87291 | -33.9804 | -1.3906
0.87901 | -33.981 | -1.4109
0.88512 | -33.9819 | -1.4312
0.89122 | -33.9828 | -1.4517
0.89733 | -33.9833 | -1.4693
0.90343 | -33.9837 | -1.4881
0.90953 | -33.9833 | -1.5023
0.91564 | -33.9833 | -1.5189
0.92174 | -33.9833 | -1.5344
0.92785 | -33.9832 | -1.5492
0.93395 | -33.9828 | -1.5635
0.94006 | -33.9826 | -1.5769
0.94616 | -33.9814 | -1.5871
0.95226 | -33.981 | -1.599
0.95837 | -33.9806 | -1.6117
0.96447 | -33.9803 | -1.6227
0.97058 | -33.9796 | -1.6341
0.97668 | -33.9791 | -1.6447
0.98279 | -33.9785 | -1.6545
0.98889 | -33.978 | -1.6629
0.99499 | -33.9776 | -1.6727
1.0011 | -33.9772 | -1.6807
1.0072 | -33.9766 | -1.6898
1.0133 | -33.9761 | -1.6973
1.0194 | -33.9757 | -1.7064
1.0255 | -33.9753 | -1.7169
1.0316 | -33.9748 | -1.7265
1.0377 | -33.9746 | -1.738
1.0438 | -33.9742 | -1.7473
1.0499 | -33.9741 | -1.7606
1.056 | -33.9736 | -1.7732
1.0621 | -33.9733 | -1.7855
1.0682 | -33.9731 | -1.7958
1.0743 | -33.9729 | -1.8101
1.0805 | -33.9729 | -1.8223
1.0866 | -33.9726 | -1.8388
1.0927 | -33.9722 | -1.8529
1.0988 | -33.972 | -1.867
1.1049 | -33.9716 | -1.8812
1.111 | -33.9713 | -1.8963
1.1171 | -33.971 | -1.9113
1.1232 | -33.971 | -1.9285
1.1293 | -33.9711 | -1.9437
1.1354 | -33.9713 | -1.9585
1.1415 | -33.9714 | -1.9753
1.1476 | -33.9711 | -1.9921
1.1537 | -33.971 | -2.0085
1.1598 | -33.9712 | -2.0246
1.1659 | -33.9713 | -2.0395
1.172 | -33.9713 | -2.0544
1.1781 | -33.9713 | -2.0691
1.1842 | -33.9714 | -2.0849
1.1903 | -33.9718 | -2.0966
1.1964 | -33.9718 | -2.1114
1.2025 | -33.9721 | -2.1261
1.2086 | -33.9725 | -2.1388
1.2147 | -33.9724 | -2.1521
1.2209 | -33.9729 | -2.1638
1.227 | -33.9733 | -2.174
1.2331 | -33.9735 | -2.1857
1.2392 | -33.9737 | -2.198
1.2453 | -33.974 | -2.2097
1.2514 | -33.9742 | -2.219
1.2575 | -33.9746 | -2.2287
1.2636 | -33.9751 | -2.2362
1.2697 | -33.9755 | -2.2456
1.2758 | -33.9761 | -2.256
1.2819 | -33.9766 | -2.2648
1.288 | -33.9769 | -2.275
1.2941 | -33.9775 | -2.2837
1.3002 | -33.9781 | -2.2917
1.3063 | -33.9784 | -2.3
1.3124 | -33.9789 | -2.3088
1.3185 | -33.9795 | -2.3168
1.3246 | -33.98 | -2.3251
1.3307 | -33.9806 | -2.3329
1.3368 | -33.9814 | -2.3421
1.3429 | -33.9819 | -2.3503
1.349 | -33.9824 | -2.3596
1.3551 | -33.9831 | -2.3689
1.3613 | -33.9835 | -2.377
1.3674 | -33.9842 | -2.3877
1.3735 | -33.9846 | -2.3982
1.3796 | -33.985 | -2.4076
1.3857 | -33.9853 | -2.4179
1.3918 | -33.9858 | -2.4283
1.3979 | -33.9864 | -2.4381
1.404 | -33.9871 | -2.4512
1.4101 | -33.9876 | -2.4625
1.4162 | -33.9881 | -2.4744
1.4223 | -33.9887 | -2.4868
1.4284 | -33.9894 | -2.4984
1.4345 | -33.9897 | -2.511
1.4406 | -33.9902 | -2.5266
1.4467 | -33.9908 | -2.5396
1.4528 | -33.9914 | -2.5524
1.4589 | -33.992 | -2.5661
1.465 | -33.9924 | -2.578
1.4711 | -33.9931 | -2.591
1.4772 | -33.9938 | -2.6056
1.4833 | -33.9944 | -2.6201
1.4894 | -33.9951 | -2.6334
1.4955 | -33.9957 | -2.6467
1.5016 | -33.9963 | -2.6604
1.5078 | -33.9967 | -2.6758
1.5139 | -33.9972 | -2.6897
1.52 | -33.9978 | -2.7034
1.5261 | -33.9984 | -2.7175
1.5322 | -33.9989 | -2.7325
1.5383 | -33.9994 | -2.7456
1.5444 | -34.0002 | -2.7592
1.5505 | -34.0007 | -2.7735
1.5566 | -34.0013 | -2.7867
1.5627 | -34.0018 | -2.8008
1.5688 | -34.0023 | -2.8147
1.5749 | -34.0028 | -2.8288
1.581 | -34.0032 | -2.8427
1.5871 | -34.0037 | -2.8569
1.5932 | -34.0042 | -2.8706
1.5993 | -34.0047 | -2.8847
1.6054 | -34.0053 | -2.8982
1.6115 | -34.0058 | -2.9131
1.6176 | -34.0062 | -2.9259
1.6237 | -34.0065 | -2.9398
1.6298 | -34.0069 | -2.9524
1.6359 | -34.0073 | -2.9657
1.642 | -34.0079 | -2.9778
1.6482 | -34.0083 | -2.991
1.6543 | -34.0089 | -3.005
1.6604 | -34.0093 | -3.0178
1.6665 | -34.0095 | -3.0308
1.6726 | -34.0098 | -3.0438
1.6787 | -34.0103 | -3.0571
1.6848 | -34.0103 | -3.0696
1.6909 | -34.0106 | -3.0823
1.697 | -34.0109 | -3.0945
1.7031 | -34.0111 | -3.1073
1.7092 | -34.0112 | -3.1195
1.7153 | -34.0115 | -3.1314
1.7214 | -34.0114 | -3.1445
1.7275 | -34.0115 | -3.1573
1.7336 | -34.0115 | -3.1705
1.7397 | -34.0117 | -3.1826
1.7458 | -34.0118 | -3.195
1.7519 | -34.0119 | -3.2076
1.758 | -34.012 | -3.2193
1.7641 | -34.0121 | -3.2306
1.7702 | -34.0121 | -3.2441
1.7763 | -34.0122 | -3.2553
1.7824 | -34.0121 | -3.2679
1.7885 | -34.012 | -3.2796
1.7947 | -34.0121 | -3.2914
1.8008 | -34.0121 | -3.3031
1.8069 | -34.0121 | -3.3143
1.813 | -34.0121 | -3.3265
1.8191 | -34.012 | -3.3388
1.8252 | -34.0119 | -3.3512
1.8313 | -34.012 | -3.3632
1.8374 | -34.0121 | -3.3748
1.8435 | -34.012 | -3.3866
1.8496 | -34.012 | -3.3988
1.8557 | -34.012 | -3.4103
1.8618 | -34.0121 | -3.4222
1.8679 | -34.0121 | -3.4339
1.874 | -34.0121 | -3.4458
1.8801 | -34.0121 | -3.4577
1.8862 | -34.0121 | -3.4697
1.8923 | -34.0121 | -3.4813
1.8984 | -34.0121 | -3.4927
1.9045 | -34.012 | -3.5039
1.9106 | -34.0119 | -3.5149
1.9167 | -34.0118 | -3.5259
1.9228 | -34.0119 | -3.5376
1.9289 | -34.0119 | -3.5481
1.9351 | -34.0119 | -3.5581
1.9412 | -34.012 | -3.5687
1.9473 | -34.0123 | -3.58
1.9534 | -34.0125 | -3.5905
1.9595 | -34.0127 | -3.6019
1.9656 | -34.013 | -3.6132
1.9717 | -34.0132 | -3.624
1.9778 | -34.0135 | -3.6344
1.9839 | -34.0137 | -3.645
1.99 | -34.014 | -3.6556
1.9961 | -34.0141 | -3.667
2.0022 | -34.0144 | -3.6781
2.0083 | -34.0145 | -3.6883
2.0144 | -34.0147 | -3.6987
2.0205 | -34.0152 | -3.7098
2.0266 | -34.0156 | -3.7204
2.0327 | -34.016 | -3.7316
2.0388 | -34.0167 | -3.7425
2.0449 | -34.0169 | -3.7531
2.051 | -34.0173 | -3.7636
2.0571 | -34.0177 | -3.7741
2.0632 | -34.0181 | -3.7845
2.0693 | -34.0184 | -3.7962
2.0754 | -34.0189 | -3.8071
2.0816 | -34.0194 | -3.8185
2.0877 | -34.0197 | -3.8296
2.0938 | -34.0203 | -3.8414
2.0999 | -34.0206 | -3.8525
2.106 | -34.0212 | -3.8636
2.1121 | -34.0217 | -3.8749
2.1182 | -34.0221 | -3.8852
2.1243 | -34.0227 | -3.8951
2.1304 | -34.0231 | -3.906
2.1365 | -34.0236 | -3.9173
2.1426 | -34.0241 | -3.9297
2.1487 | -34.0247 | -3.9406
2.1548 | -34.0253 | -3.952
2.1609 | -34.0261 | -3.9629
2.167 | -34.0267 | -3.9742
2.1731 | -34.0272 | -3.9859
2.1792 | -34.0276 | -3.9973
2.1853 | -34.0281 | -4.0088
2.1914 | -34.0286 | -4.0201
2.1975 | -34.0292 | -4.0318
2.2036 | -34.0295 | -4.0433
2.2097 | -34.0299 | -4.0545
2.2158 | -34.0304 | -4.0656
2.222 | -34.0309 | -4.076
2.2281 | -34.0314 | -4.0875
2.2342 | -34.0318 | -4.0988
2.2403 | -34.0322 | -4.1096
2.2464 | -34.0324 | -4.1207
2.2525 | -34.0328 | -4.1315
2.2586 | -34.0332 | -4.1421
2.2647 | -34.0335 | -4.1528
2.2708 | -34.034 | -4.1644
2.2769 | -34.0343 | -4.1748
2.283 | -34.0344 | -4.1863
2.2891 | -34.0346 | -4.1973
2.2952 | -34.0347 | -4.208
2.3013 | -34.0349 | -4.2191
2.3074 | -34.0352 | -4.2299
2.3135 | -34.0354 | -4.2411
2.3196 | -34.0353 | -4.2517
2.3257 | -34.0354 | -4.2621
2.3318 | -34.0356 | -4.2738
2.3379 | -34.0357 | -4.2838
2.344 | -34.0358 | -4.2956
2.3501 | -34.0359 | -4.3067
2.3562 | -34.0357 | -4.3168
2.3623 | -34.0357 | -4.3275
2.3685 | -34.0356 | -4.3383
2.3746 | -34.0356 | -4.3496
2.3807 | -34.0356 | -4.3604
2.3868 | -34.0355 | -4.3709
2.3929 | -34.0355 | -4.3816
2.399 | -34.0353 | -4.392
2.4051 | -34.0354 | -4.4026
2.4112 | -34.0354 | -4.4138
2.4173 | -34.0353 | -4.4248
2.4234 | -34.035 | -4.4358
2.4295 | -34.0349 | -4.4469
2.4356 | -34.0348 | -4.4582
2.4417 | -34.0346 | -4.4696
2.4478 | -34.0344 | -4.4808
2.4539 | -34.0343 | -4.4912
2.46 | -34.0341 | -4.5028
2.4661 | -34.034 | -4.5138
2.4722 | -34.0338 | -4.5244
2.4783 | -34.0336 | -4.5353
2.4844 | -34.0334 | -4.5467
2.4905 | -34.0333 | -4.5574
2.4966 | -34.0332 | -4.5688
2.5027 | -34.033 | -4.5792
2.5089 | -34.0329 | -4.5907
2.515 | -34.0326 | -4.6012
2.5211 | -34.0324 | -4.6124
2.5272 | -34.0322 | -4.6234
2.5333 | -34.032 | -4.6344
2.5394 | -34.0319 | -4.6454
2.5455 | -34.0317 | -4.6567
2.5516 | -34.0316 | -4.6674
2.5577 | -34.0315 | -4.679
2.5638 | -34.0315 | -4.6908
2.5699 | -34.0313 | -4.7024
2.576 | -34.0311 | -4.714
2.5821 | -34.031 | -4.7253
2.5882 | -34.0309 | -4.7369
2.5943 | -34.0308 | -4.7477
2.6004 | -34.0309 | -4.7591
2.6065 | -34.0309 | -4.7697
2.6126 | -34.0308 | -4.7801
2.6187 | -34.0307 | -4.7906
2.6248 | -34.0307 | -4.8015
2.6309 | -34.0307 | -4.8134
2.637 | -34.0309 | -4.8249
2.6431 | -34.0309 | -4.8353
2.6492 | -34.0309 | -4.8471
2.6554 | -34.0309 | -4.8579
2.6615 | -34.0309 | -4.8688
2.6676 | -34.031 | -4.8805
2.6737 | -34.0313 | -4.8918
2.6798 | -34.0315 | -4.9029
2.6859 | -34.0316 | -4.9141
2.692 | -34.0316 | -4.9247
2.6981 | -34.0317 | -4.9358
2.7042 | -34.032 | -4.9465
2.7103 | -34.032 | -4.9572
2.7164 | -34.0323 | -4.9681
2.7225 | -34.0325 | -4.9796
2.7286 | -34.0327 | -4.9907
2.7347 | -34.0328 | -5.0016
2.7408 | -34.0332 | -5.0115
2.7469 | -34.0334 | -5.0221
2.753 | -34.0336 | -5.0326
2.7591 | -34.0339 | -5.0438
2.7652 | -34.0341 | -5.0553
2.7713 | -34.0343 | -5.066
2.7774 | -34.0347 | -5.0764
2.7835 | -34.0349 | -5.0874
2.7896 | -34.035 | -5.0977
2.7958 | -34.0353 | -5.1083
2.8019 | -34.0356 | -5.1184
2.808 | -34.0359 | -5.1293
2.8141 | -34.0363 | -5.1403
2.8202 | -34.0366 | -5.1511
2.8263 | -34.0369 | -5.1617
2.8324 | -34.0374 | -5.1729
2.8385 | -34.0378 | -5.1834
2.8446 | -34.0382 | -5.1939
2.8507 | -34.0386 | -5.2044
2.8568 | -34.039 | -5.2144
2.8629 | -34.0394 | -5.2245
2.869 | -34.0397 | -5.2354
2.8751 | -34.0401 | -5.246
2.8812 | -34.0407 | -5.2569
2.8873 | -34.041 | -5.2674
2.8934 | -34.0414 | -5.2779
2.8995 | -34.0418 | -5.2885
2.9056 | -34.0422 | -5.2997
2.9117 | -34.0426 | -5.3106
2.9178 | -34.043 | -5.321
2.9239 | -34.0435 | -5.3323
2.93 | -34.044 | -5.3435
2.9361 | -34.0444 | -5.3541
2.9423 | -34.0448 | -5.3647
2.9484 | -34.0453 | -5.3752
2.9545 | -34.0458 | -5.3857
2.9606 | -34.0463 | -5.3966
2.9667 | -34.0467 | -5.4078
2.9728 | -34.0472 | -5.4191
2.9789 | -34.0477 | -5.4297
2.985 | -34.0482 | -5.4405
2.9911 | -34.0486 | -5.4517
2.9972 | -34.0492 | -5.4624
3.0033 | -34.0496 | -5.4737
3.0094 | -34.05 | -5.4844
3.0155 | -34.0505 | -5.4953
3.0216 | -34.0509 | -5.5063
3.0277 | -34.0513 | -5.5174
3.0338 | -34.0517 | -5.5284
3.0399 | -34.0522 | -5.5392
3.046 | -34.0526 | -5.5504
3.0521 | -34.0531 | -5.5612
3.0582 | -34.0535 | -5.573
3.0643 | -34.0539 | -5.5839
3.0704 | -34.0543 | -5.5955
3.0765 | -34.0546 | -5.6062
3.0827 | -34.055 | -5.6172
3.0888 | -34.0554 | -5.6282
3.0949 | -34.0559 | -5.6395
3.101 | -34.0563 | -5.6507
3.1071 | -34.0567 | -5.6624
3.1132 | -34.0572 | -5.674
3.1193 | -34.0576 | -5.6854
3.1254 | -34.058 | -5.6966
3.1315 | -34.0585 | -5.7085
3.1376 | -34.0588 | -5.7197
3.1437 | -34.0592 | -5.7313
3.1498 | -34.0596 | -5.7433
3.1559 | -34.0599 | -5.7544
3.162 | -34.0602 | -5.7663
3.1681 | -34.0605 | -5.7779
3.1742 | -34.0609 | -5.7896
3.1803 | -34.0612 | -5.8013
3.1864 | -34.0616 | -5.8125
3.1925 | -34.0619 | -5.8238
3.1986 | -34.0622 | -5.8359
3.2047 | -34.0626 | -5.8477
3.2108 | -34.0629 | -5.859
3.2169 | -34.0633 | -5.8704
3.223 | -34.0635 | -5.882
3.2292 | -34.0638 | -5.8935
3.2353 | -34.0641 | -5.9045
3.2414 | -34.0643 | -5.9156
3.2475 | -34.0646 | -5.9275
3.2536 | -34.0649 | -5.9386
3.2597 | -34.0651 | -5.9496
3.2658 | -34.0654 | -5.9607
3.2719 | -34.0656 | -5.9719
3.278 | -34.0659 | -5.9834
3.2841 | -34.0661 | -5.9942
3.2902 | -34.0664 | -6.0054
3.2963 | -34.0666 | -6.0168
3.3024 | -34.0667 | -6.0281
3.3085 | -34.067 | -6.0394
3.3146 | -34.0672 | -6.0507
3.3207 | -34.0674 | -6.0623
3.3268 | -34.0676 | -6.0737
3.3329 | -34.0677 | -6.0848
3.339 | -34.0679 | -6.0957
3.3451 | -34.068 | -6.1066
3.3512 | -34.0681 | -6.1173
3.3573 | -34.0684 | -6.1284
3.3634 | -34.0685 | -6.14
3.3696 | -34.0687 | -6.1511
3.3757 | -34.0687 | -6.1618
3.3818 | -34.0689 | -6.1724
3.3879 | -34.0691 | -6.1835
3.394 | -34.0692 | -6.1948
3.4001 | -34.0693 | -6.2058
3.4062 | -34.0694 | -6.2167
3.4123 | -34.0695 | -6.2274
3.4184 | -34.0697 | -6.2389
3.4245 | -34.0699 | -6.2498
3.4306 | -34.07 | -6.2605
3.4367 | -34.0701 | -6.2708
3.4428 | -34.0702 | -6.2813
3.4489 | -34.0703 | -6.2923
3.455 | -34.0704 | -6.3033
3.4611 | -34.0706 | -6.3144
3.4672 | -34.0708 | -6.3252
3.4733 | -34.0709 | -6.3362
3.4794 | -34.0711 | -6.3469
3.4855 | -34.0712 | -6.3571
3.4916 | -34.0714 | -6.3675
3.4977 | -34.0715 | -6.3786
3.5038 | -34.0715 | -6.3895
3.5099 | -34.0716 | -6.3998
3.5161 | -34.0717 | -6.4102
3.5222 | -34.0718 | -6.4205
3.5283 | -34.0718 | -6.4311
3.5344 | -34.0719 | -6.4417
3.5405 | -34.072 | -6.4522
3.5466 | -34.0721 | -6.4634
3.5527 | -34.0722 | -6.4741
3.5588 | -34.0722 | -6.4847
3.5649 | -34.0722 | -6.4951
3.571 | -34.0723 | -6.5058
3.5771 | -34.0723 | -6.5169
3.5832 | -34.0724 | -6.5264
3.5893 | -34.0726 | -6.5372
3.5954 | -34.0728 | -6.5479
3.6015 | -34.0729 | -6.5587
3.6076 | -34.0731 | -6.5692
3.6137 | -34.0732 | -6.5793
3.6198 | -34.0734 | -6.5895
3.6259 | -34.0735 | -6.6004
3.632 | -34.0737 | -6.6112
3.6381 | -34.0738 | -6.6217
3.6442 | -34.0739 | -6.6323
3.6503 | -34.0741 | -6.6422
3.6565 | -34.0743 | -6.653
3.6626 | -34.0743 | -6.6637
3.6687 | -34.0745 | -6.6741
3.6748 | -34.0747 | -6.6843
3.6809 | -34.0749 | -6.695
3.687 | -34.0751 | -6.7057
3.6931 | -34.0753 | -6.7159
3.6992 | -34.0754 | -6.7265
3.7053 | -34.0756 | -6.7369
3.7114 | -34.0756 | -6.7471
3.7175 | -34.0758 | -6.7579
3.7236 | -34.076 | -6.768
3.7297 | -34.0762 | -6.7789
3.7358 | -34.0764 | -6.7897
3.7419 | -34.0766 | -6.8005
3.748 | -34.0768 | -6.8107
3.7541 | -34.077 | -6.8213
3.7602 | -34.0773 | -6.8319
3.7663 | -34.0774 | -6.8423
3.7724 | -34.0776 | -6.8531
3.7785 | -34.0779 | -6.8635
3.7846 | -34.0781 | -6.8738
3.7907 | -34.0783 | -6.8847
3.7969 | -34.0786 | -6.8954
3.803 | -34.0788 | -6.9056
3.8091 | -34.079 | -6.9163
3.8152 | -34.0793 | -6.9267
3.8213 | -34.0795 | -6.9377
3.8274 | -34.0797 | -6.948
3.8335 | -34.0801 | -6.9589
3.8396 | -34.0804 | -6.9696
3.8457 | -34.0807 | -6.9806
3.8518 | -34.081 | -6.9909
3.8579 | -34.0813 | -7.0016
3.864 | -34.0816 | -7.0125
3.8701 | -34.0819 | -7.0234
3.8762 | -34.0821 | -7.0338
3.8823 | -34.0824 | -7.0445
3.8884 | -34.0827 | -7.0554
3.8945 | -34.0831 | -7.066
3.9006 | -34.0834 | -7.0772
3.9067 | -34.0837 | -7.0881
3.9128 | -34.084 | -7.099
3.9189 | -34.0843 | -7.1098
3.925 | -34.0846 | -7.1206
3.9311 | -34.0848 | -7.1309
3.9372 | -34.0852 | -7.1417
3.9434 | -34.0854 | -7.1524
3.9495 | -34.0858 | -7.1632
3.9556 | -34.0861 | -7.1736
3.9617 | -34.0864 | -7.1841
3.9678 | -34.0868 | -7.1946
3.9739 | -34.0871 | -7.2051
3.98 | -34.0874 | -7.2161
3.9861 | -34.0877 | -7.2266
3.9922 | -34.0879 | -7.2372
3.9983 | -34.0882 | -7.2474
4.0044 | -34.0885 | -7.2576
4.0105 | -34.0889 | -7.2685
4.0166 | -34.0892 | -7.2792
4.0227 | -34.0896 | -7.2896
4.0288 | -34.0899 | -7.3008
4.0349 | -34.0901 | -7.3114
4.041 | -34.0904 | -7.3221
4.0471 | -34.0906 | -7.3328
4.0532 | -34.091 | -7.343
4.0593 | -34.0914 | -7.3539
4.0654 | -34.0918 | -7.3644
4.0715 | -34.0921 | -7.3758
4.0776 | -34.0924 | -7.3861
4.0838 | -34.0927 | -7.3972
4.0899 | -34.0931 | -7.4077
4.096 | -34.0934 | -7.4184
4.1021 | -34.0936 | -7.4286
4.1082 | -34.0939 | -7.4392
4.1143 | -34.0942 | -7.4494
4.1204 | -34.0945 | -7.4596
4.1265 | -34.0949 | -7.47
4.1326 | -34.0952 | -7.4803
4.1387 | -34.0956 | -7.4906
4.1448 | -34.0959 | -7.5013
4.1509 | -34.0962 | -7.5119
4.157 | -34.0965 | -7.5228
4.1631 | -34.0968 | -7.5337
4.1692 | -34.0971 | -7.5441
4.1753 | -34.0974 | -7.5546
4.1814 | -34.0976 | -7.5649
4.1875 | -34.098 | -7.5756
4.1936 | -34.0983 | -7.5861
4.1997 | -34.0985 | -7.5964
4.2058 | -34.0988 | -7.6072
4.2119 | -34.0991 | -7.6177
4.218 | -34.0994 | -7.6279
4.2241 | -34.0996 | -7.6381
4.2303 | -34.1 | -7.648
4.2364 | -34.1003 | -7.6584
4.2425 | -34.1005 | -7.6692
4.2486 | -34.1008 | -7.6796
4.2547 | -34.101 | -7.6897
4.2608 | -34.1013 | -7.7003
4.2669 | -34.1015 | -7.7104
4.273 | -34.1018 | -7.721
4.2791 | -34.1021 | -7.7314
4.2852 | -34.1024 | -7.7422
4.2913 | -34.1026 | -7.7524
4.2974 | -34.1029 | -7.763
4.3035 | -34.1032 | -7.7739
4.3096 | -34.1034 | -7.7843
4.3157 | -34.1037 | -7.7947
4.3218 | -34.1039 | -7.8049
4.3279 | -34.1042 | -7.8154
4.334 | -34.1044 | -7.8258
4.3401 | -34.1047 | -7.8358
4.3462 | -34.105 | -7.8462
4.3523 | -34.1052 | -7.8561
4.3584 | -34.1054 | -7.8661
4.3645 | -34.1057 | -7.8763
4.3707 | -34.1059 | -7.8864
4.3768 | -34.1062 | -7.8967
4.3829 | -34.1065 | -7.9074
4.389 | -34.1067 | -7.9176
4.3951 | -34.1069 | -7.9277
4.4012 | -34.1071 | -7.9384
4.4073 | -34.1073 | -7.9483
4.4134 | -34.1076 | -7.9585
4.4195 | -34.1079 | -7.9688
4.4256 | -34.1082 | -7.9787
4.4317 | -34.1084 | -7.9888
4.4378 | -34.1086 | -7.9986
4.4439 | -34.1088 | -8.0086
4.45 | -34.1091 | -8.0188
4.4561 | -34.1093 | -8.0292
4.4622 | -34.1095 | -8.0394
4.4683 | -34.1097 | -8.0495
4.4744 | -34.11 | -8.0601
4.4805 | -34.1102 | -8.07
4.4866 | -34.1105 | -8.0803
4.4927 | -34.1107 | -8.0907
4.4988 | -34.1109 | -8.1008
4.5049 | -34.1112 | -8.1106
4.511 | -34.1114 | -8.1206
4.5172 | -34.1116 | -8.1308
4.5233 | -34.1118 | -8.1412
4.5294 | -34.1121 | -8.1512
4.5355 | -34.1123 | -8.1614
4.5416 | -34.1125 | -8.1717
4.5477 | -34.1128 | -8.1819
4.5538 | -34.113 | -8.1917
4.5599 | -34.1132 | -8.2016
4.566 | -34.1134 | -8.2121
4.5721 | -34.1136 | -8.2222
4.5782 | -34.1139 | -8.2324
4.5843 | -34.1141 | -8.2429
4.5904 | -34.1143 | -8.2531
4.5965 | -34.1145 | -8.2632
4.6026 | -34.1148 | -8.2738
4.6087 | -34.1149 | -8.2839
4.6148 | -34.1151 | -8.2944
4.6209 | -34.1153 | -8.3046
4.627 | -34.1156 | -8.3149
4.6331 | -34.1158 | -8.3251
4.6392 | -34.116 | -8.3353
4.6453 | -34.1162 | -8.3455
4.6514 | -34.1164 | -8.3558
4.6576 | -34.1166 | -8.3664
4.6637 | -34.1167 | -8.3761
4.6698 | -34.117 | -8.3861
4.6759 | -34.1171 | -8.3966
4.682 | -34.1173 | -8.407
4.6881 | -34.1175 | -8.4172
4.6942 | -34.1177 | -8.4273
4.7003 | -34.1178 | -8.4374
4.7064 | -34.118 | -8.4475
4.7125 | -34.1182 | -8.458
4.7186 | -34.1184 | -8.4678
4.7247 | -34.1186 | -8.4781
4.7308 | -34.1188 | -8.4884
4.7369 | -34.119 | -8.4988
4.743 | -34.1193 | -8.5093
4.7491 | -34.1194 | -8.5198
4.7552 | -34.1196 | -8.53
4.7613 | -34.1198 | -8.5405
4.7674 | -34.12 | -8.5506
4.7735 | -34.1201 | -8.5613
4.7796 | -34.1204 | -8.5719
4.7857 | -34.1206 | -8.5823
4.7918 | -34.1208 | -8.5922
4.7979 | -34.1211 | -8.6023
4.8041 | -34.1213 | -8.6126
4.8102 | -34.1215 | -8.6225
4.8163 | -34.1217 | -8.6326
4.8224 | -34.1219 | -8.6431
4.8285 | -34.1222 | -8.6531
4.8346 | -34.1224 | -8.6631
4.8407 | -34.1225 | -8.6735
4.8468 | -34.1227 | -8.6841
4.8529 | -34.1229 | -8.6947
4.859 | -34.1231 | -8.7053
4.8651 | -34.1233 | -8.7153
4.8712 | -34.1235 | -8.7254
4.8773 | -34.1236 | -8.7357
4.8834 | -34.1238 | -8.7461
4.8895 | -34.1241 | -8.7562
4.8956 | -34.1242 | -8.7665
4.9017 | -34.1244 | -8.7765
4.9078 | -34.1246 | -8.7872
4.9139 | -34.1248 | -8.7971
4.92 | -34.125 | -8.8071
4.9261 | -34.1253 | -8.8168
4.9322 | -34.1254 | -8.8273
4.9383 | -34.1257 | -8.8376
4.9445 | -34.1258 | -8.8478
4.9506 | -34.126 | -8.8581
4.9567 | -34.1263 | -8.8684
4.9628 | -34.1265 | -8.8785
4.9689 | -34.1267 | -8.8892
4.975 | -34.1269 | -8.8999
4.9811 | -34.1271 | -8.9097
4.9872 | -34.1274 | -8.9196
4.9933 | -34.1277 | -8.9303
4.9994 | -34.128 | -8.9411
5.0055 | -34.1283 | -8.9513
5.0116 | -34.1285 | -8.962
5.0177 | -34.1288 | -8.9725
5.0238 | -34.1291 | -8.9833
5.0299 | -34.1294 | -8.9938
5.036 | -34.1297 | -9.0042
5.0421 | -34.1298 | -9.0145
5.0482 | -34.1301 | -9.0249
5.0543 | -34.1303 | -9.0348
5.0604 | -34.1306 | -9.0452
5.0665 | -34.1309 | -9.0553
5.0726 | -34.1311 | -9.0658
5.0787 | -34.1314 | -9.0764
5.0848 | -34.1317 | -9.0868
5.091 | -34.1319 | -9.0969
5.0971 | -34.1323 | -9.1069
5.1032 | -34.1325 | -9.1177
5.1093 | -34.1328 | -9.1281
5.1154 | -34.133 | -9.1381
5.1215 | -34.1332 | -9.1482
5.1276 | -34.1335 | -9.1585
5.1337 | -34.1339 | -9.1687
5.1398 | -34.1341 | -9.1788
5.1459 | -34.1344 | -9.1895
5.152 | -34.1347 | -9.2
5.1581 | -34.135 | -9.2105
5.1642 | -34.1353 | -9.2207
5.1703 | -34.1356 | -9.231
5.1764 | -34.1359 | -9.2412
5.1825 | -34.1362 | -9.2515
5.1886 | -34.1365 | -9.2614
5.1947 | -34.1368 | -9.2718
5.2008 | -34.1371 | -9.2821
5.2069 | -34.1374 | -9.2926
5.213 | -34.1377 | -9.3027
5.2191 | -34.1381 | -9.3131
5.2252 | -34.1384 | -9.3235
5.2314 | -34.1386 | -9.3341
5.2375 | -34.139 | -9.3443
5.2436 | -34.1393 | -9.3547
5.2497 | -34.1396 | -9.3646
5.2558 | -34.14 | -9.3753
5.2619 | -34.1403 | -9.386
5.268 | -34.1407 | -9.3966
5.2741 | -34.1411 | -9.4071
5.2802 | -34.1414 | -9.4178
5.2863 | -34.1418 | -9.4283
5.2924 | -34.1421 | -9.4388
5.2985 | -34.1425 | -9.4492
5.3046 | -34.1428 | -9.4596
5.3107 | -34.1432 | -9.4701
5.3168 | -34.1435 | -9.4806
5.3229 | -34.1439 | -9.4909
5.329 | -34.1442 | -9.5014
5.3351 | -34.1445 | -9.5114
5.3412 | -34.1449 | -9.5217
5.3473 | -34.1452 | -9.5323
5.3534 | -34.1456 | -9.5429
5.3595 | -34.1459 | -9.5533
5.3656 | -34.1464 | -9.5636
5.3717 | -34.1467 | -9.5744
5.3779 | -34.1471 | -9.5854
5.384 | -34.1474 | -9.5956
5.3901 | -34.1478 | -9.6064
5.3962 | -34.1482 | -9.6165
5.4023 | -34.1486 | -9.6272
5.4084 | -34.149 | -9.6377
5.4145 | -34.1494 | -9.6479
5.4206 | -34.1497 | -9.6582
5.4267 | -34.1501 | -9.6691
5.4328 | -34.1505 | -9.6795
5.4389 | -34.1508 | -9.6896
5.445 | -34.1512 | -9.7005
5.4511 | -34.1516 | -9.7107
5.4572 | -34.1519 | -9.7213
5.4633 | -34.1523 | -9.7317
5.4694 | -34.1526 | -9.7421
5.4755 | -34.153 | -9.7522
5.4816 | -34.1534 | -9.7629
5.4877 | -34.1538 | -9.7735
5.4938 | -34.1542 | -9.7836
5.4999 | -34.1546 | -9.7939
5.506 | -34.155 | -9.8044
5.5121 | -34.1553 | -9.8147
5.5183 | -34.1557 | -9.825
5.5244 | -34.156 | -9.8356
5.5305 | -34.1565 | -9.846
5.5366 | -34.1568 | -9.8561
5.5427 | -34.1571 | -9.8669
5.5488 | -34.1575 | -9.8774
5.5549 | -34.1579 | -9.888
5.561 | -34.1582 | -9.8985
5.5671 | -34.1586 | -9.9088
5.5732 | -34.1589 | -9.9195
5.5793 | -34.1593 | -9.9302
5.5854 | -34.1597 | -9.9406
5.5915 | -34.1601 | -9.9514
5.5976 | -34.1605 | -9.962
5.6037 | -34.1608 | -9.9725
5.6098 | -34.1612 | -9.983
5.6159 | -34.1616 | -9.9932
5.622 | -34.162 | -10.0037
5.6281 | -34.1624 | -10.0143
5.6342 | -34.1627 | -10.0251
5.6403 | -34.1631 | -10.0356
5.6464 | -34.1634 | -10.0462
5.6525 | -34.1638 | -10.0568
5.6586 | -34.1641 | -10.067
5.6648 | -34.1645 | -10.0777
5.6709 | -34.165 | -10.0882
5.677 | -34.1654 | -10.0987
5.6831 | -34.1657 | -10.1093
5.6892 | -34.166 | -10.12
5.6953 | -34.1664 | -10.1304
5.7014 | -34.1668 | -10.1411
5.7075 | -34.1671 | -10.1516
5.7136 | -34.1674 | -10.162
5.7197 | -34.1677 | -10.1725
5.7258 | -34.1681 | -10.1829
5.7319 | -34.1684 | -10.1934
5.738 | -34.1688 | -10.2041
5.7441 | -34.1691 | -10.2144
5.7502 | -34.1695 | -10.2247
5.7563 | -34.1699 | -10.2351
5.7624 | -34.1703 | -10.2459
5.7685 | -34.1706 | -10.2565
5.7746 | -34.1709 | -10.2672
5.7807 | -34.1712 | -10.2778
5.7868 | -34.1715 | -10.2885
5.7929 | -34.1719 | -10.2988
5.799 | -34.1722 | -10.3091
5.8052 | -34.1725 | -10.3195
5.8113 | -34.1729 | -10.3301
5.8174 | -34.1732 | -10.3404
5.8235 | -34.1736 | -10.3509
5.8296 | -34.1739 | -10.3615
5.8357 | -34.1742 | -10.3719
5.8418 | -34.1746 | -10.3825
5.8479 | -34.1749 | -10.393
5.854 | -34.1752 | -10.4035
5.8601 | -34.1755 | -10.414
5.8662 | -34.1758 | -10.4246
5.8723 | -34.1761 | -10.4348
5.8784 | -34.1764 | -10.4452
5.8845 | -34.1767 | -10.4556
5.8906 | -34.177 | -10.4663
5.8967 | -34.1773 | -10.4765
5.9028 | -34.1777 | -10.4867
5.9089 | -34.1781 | -10.4969
5.915 | -34.1783 | -10.5073
5.9211 | -34.1786 | -10.5176
5.9272 | -34.179 | -10.5277
5.9333 | -34.1793 | -10.5381
5.9394 | -34.1796 | -10.5487
5.9455 | -34.1799 | -10.5589
5.9517 | -34.1803 | -10.569
5.9578 | -34.1806 | -10.5797
5.9639 | -34.181 | -10.5897
5.97 | -34.1813 | -10.6002
5.9761 | -34.1816 | -10.6104
5.9822 | -34.1819 | -10.6209
5.9883 | -34.1822 | -10.6315
5.9944 | -34.1825 | -10.6419
6.0005 | -34.1828 | -10.6526
6.0066 | -34.1831 | -10.6627
6.0127 | -34.1834 | -10.6727
6.0188 | -34.1837 | -10.6834
6.0249 | -34.184 | -10.694
6.031 | -34.1843 | -10.7042
6.0371 | -34.1846 | -10.7142
6.0432 | -34.1849 | -10.7245
6.0493 | -34.1852 | -10.7347
6.0554 | -34.1855 | -10.7449
6.0615 | -34.1858 | -10.7554
6.0676 | -34.1861 | -10.766
6.0737 | -34.1864 | -10.7764
6.0798 | -34.1867 | -10.7868
6.0859 | -34.187 | -10.7971
6.0921 | -34.1873 | -10.8071
6.0982 | -34.1876 | -10.8173
6.1043 | -34.1879 | -10.8275
6.1104 | -34.1882 | -10.838
6.1165 | -34.1885 | -10.8479
6.1226 | -34.1888 | -10.8582
6.1287 | -34.1891 | -10.8685
6.1348 | -34.1894 | -10.8786
6.1409 | -34.1896 | -10.8889
6.147 | -34.1899 | -10.8991
6.1531 | -34.1902 | -10.9093
6.1592 | -34.1905 | -10.9195
6.1653 | -34.1908 | -10.9296
6.1714 | -34.191 | -10.94
6.1775 | -34.1914 | -10.9504
6.1836 | -34.1916 | -10.9607
6.1897 | -34.1919 | -10.9709
6.1958 | -34.1922 | -10.9814
6.2019 | -34.1925 | -10.9919
6.208 | -34.1928 | -11.002
6.2141 | -34.1931 | -11.0123
6.2202 | -34.1934 | -11.0226
6.2263 | -34.1937 | -11.0328
6.2325 | -34.194 | -11.0428
6.2386 | -34.1943 | -11.0527
6.2447 | -34.1946 | -11.0633
6.2508 | -34.1949 | -11.0734
6.2569 | -34.1952 | -11.0834
6.263 | -34.1954 | -11.0937
6.2691 | -34.1957 | -11.1037
6.2752 | -34.196 | -11.114
6.2813 | -34.1962 | -11.1242
6.2874 | -34.1965 | -11.1344
6.2935 | -34.1968 | -11.1445
6.2996 | -34.197 | -11.1548
6.3057 | -34.1973 | -11.165
6.3118 | -34.1976 | -11.1752
6.3179 | -34.1978 | -11.1854
6.324 | -34.1981 | -11.1957
6.3301 | -34.1984 | -11.2059
6.3362 | -34.1986 | -11.2159
6.3423 | -34.1989 | -11.2262
6.3484 | -34.1992 | -11.2365
6.3545 | -34.1995 | -11.2465
6.3606 | -34.1997 | -11.2565
6.3667 | -34.2 | -11.2664
6.3728 | -34.2003 | -11.2766
6.379 | -34.2005 | -11.2867
6.3851 | -34.2008 | -11.2967
6.3912 | -34.2011 | -11.3069
6.3973 | -34.2013 | -11.3171
6.4034 | -34.2016 | -11.3271
6.4095 | -34.2019 | -11.3372
6.4156 | -34.2021 | -11.3475
6.4217 | -34.2024 | -11.3576
6.4278 | -34.2027 | -11.3679
6.4339 | -34.203 | -11.3778
6.44 | -34.2033 | -11.3878
6.4461 | -34.2035 | -11.3981
6.4522 | -34.2038 | -11.4081
6.4583 | -34.204 | -11.4183
6.4644 | -34.2043 | -11.4284
6.4705 | -34.2046 | -11.4383
6.4766 | -34.2049 | -11.4486
6.4827 | -34.2051 | -11.4586
6.4888 | -34.2054 | -11.4687
6.4949 | -34.2057 | -11.4783
6.501 | -34.2059 | -11.4887
6.5071 | -34.2062 | -11.4987
6.5132 | -34.2065 | -11.5088
6.5194 | -34.2068 | -11.519
6.5255 | -34.207 | -11.529
6.5316 | -34.2073 | -11.5391
6.5377 | -34.2075 | -11.5493
6.5438 | -34.2078 | -11.5592
6.5499 | -34.208 | -11.5692
6.556 | -34.2083 | -11.5793
6.5621 | -34.2086 | -11.5897
6.5682 | -34.2089 | -11.5999
6.5743 | -34.2092 | -11.6098
6.5804 | -34.2095 | -11.6201
6.5865 | -34.2097 | -11.6302
6.5926 | -34.21 | -11.6401
6.5987 | -34.2103 | -11.65
6.6048 | -34.2106 | -11.6602
6.6109 | -34.2108 | -11.6704
6.617 | -34.2111 | -11.6806
6.6231 | -34.2113 | -11.6908
6.6292 | -34.2116 | -11.7009
6.6353 | -34.2119 | -11.7114
6.6414 | -34.2121 | -11.7214
6.6475 | -34.2124 | -11.7316
6.6536 | -34.2127 | -11.7417
6.6597 | -34.213 | -11.7516
6.6659 | -34.2132 | -11.7617
6.672 | -34.2135 | -11.7718
6.6781 | -34.2138 | -11.7819
6.6842 | -34.2141 | -11.7918
6.6903 | -34.2144 | -11.8019
6.6964 | -34.2147 | -11.812
6.7025 | -34.215 | -11.8219
6.7086 | -34.2152 | -11.832
6.7147 | -34.2155 | -11.8422
6.7208 | -34.2158 | -11.8526
6.7269 | -34.216 | -11.8626
6.733 | -34.2163 | -11.8724
6.7391 | -34.2166 | -11.8824
6.7452 | -34.2169 | -11.8922
6.7513 | -34.2171 | -11.9021
6.7574 | -34.2174 | -11.9121
6.7635 | -34.2177 | -11.9221
6.7696 | -34.218 | -11.9322
6.7757 | -34.2183 | -11.9422
6.7818 | -34.2186 | -11.9522
6.7879 | -34.2189 | -11.9622
6.794 | -34.2192 | -11.9721
6.8001 | -34.2194 | -11.9822
6.8063 | -34.2197 | -11.9921
6.8124 | -34.2199 | -12.0025
6.8185 | -34.2202 | -12.0127
6.8246 | -34.2205 | -12.0227
6.8307 | -34.2207 | -12.0327
6.8368 | -34.221 | -12.0431
6.8429 | -34.2213 | -12.0532
6.849 | -34.2215 | -12.0634
6.8551 | -34.2219 | -12.0735
6.8612 | -34.2221 | -12.0834
6.8673 | -34.2224 | -12.0933
6.8734 | -34.2226 | -12.1035
6.8795 | -34.2229 | -12.1137
6.8856 | -34.2232 | -12.1237
6.8917 | -34.2235 | -12.1339
6.8978 | -34.2238 | -12.1438
6.9039 | -34.2241 | -12.1537
6.91 | -34.2244 | -12.1639
6.9161 | -34.2247 | -12.1738
6.9222 | -34.225 | -12.1836
6.9283 | -34.2253 | -12.1934
6.9344 | -34.2257 | -12.2033
6.9405 | -34.226 | -12.2133
6.9466 | -34.2263 | -12.2235
6.9528 | -34.2266 | -12.2338
6.9589 | -34.227 | -12.2438
6.965 | -34.2273 | -12.2542
6.9711 | -34.2276 | -12.2644
6.9772 | -34.2279 | -12.2743
6.9833 | -34.2282 | -12.2844
6.9894 | -34.2285 | -12.2948
6.9955 | -34.2288 | -12.3048
7.0016 | -34.2291 | -12.3147
7.0077 | -34.2295 | -12.3247
7.0138 | -34.2298 | -12.3348
7.0199 | -34.2301 | -12.3446
7.026 | -34.2304 | -12.3546
7.0321 | -34.2307 | -12.3643
7.0382 | -34.231 | -12.3741
7.0443 | -34.2313 | -12.3842
7.0504 | -34.2316 | -12.3942
7.0565 | -34.232 | -12.4043
7.0626 | -34.2323 | -12.414
7.0687 | -34.2326 | -12.4237
7.0748 | -34.2329 | -12.4336
7.0809 | -34.2332 | -12.4437
7.087 | -34.2336 | -12.4539
7.0932 | -34.2339 | -12.464
7.0993 | -34.2342 | -12.4741
7.1054 | -34.2345 | -12.4839
7.1115 | -34.2349 | -12.4936
7.1176 | -34.2351 | -12.5035
7.1237 | -34.2354 | -12.5136
7.1298 | -34.2357 | -12.5233
7.1359 | -34.236 | -12.5334
7.142 | -34.2364 | -12.5432
7.1481 | -34.2367 | -12.553
7.1542 | -34.237 | -12.563
7.1603 | -34.2373 | -12.5731
7.1664 | -34.2377 | -12.5829
7.1725 | -34.2381 | -12.5929
7.1786 | -34.2384 | -12.6028
7.1847 | -34.2387 | -12.6129
7.1908 | -34.2391 | -12.6231
7.1969 | -34.2394 | -12.6331
7.203 | -34.2397 | -12.643
7.2091 | -34.24 | -12.653
7.2152 | -34.2404 | -12.6628
7.2213 | -34.2408 | -12.6727
7.2274 | -34.2411 | -12.6826
7.2335 | -34.2414 | -12.6924
7.2397 | -34.2418 | -12.7022
7.2458 | -34.2422 | -12.712
7.2519 | -34.2425 | -12.7219
7.258 | -34.2428 | -12.7321
7.2641 | -34.2432 | -12.742
7.2702 | -34.2435 | -12.7521
7.2763 | -34.2438 | -12.7621
7.2824 | -34.2442 | -12.7723
7.2885 | -34.2446 | -12.7822
7.2946 | -34.2449 | -12.7922
7.3007 | -34.2452 | -12.8022
7.3068 | -34.2455 | -12.8125
7.3129 | -34.2459 | -12.8225
7.319 | -34.2462 | -12.8325
7.3251 | -34.2465 | -12.8424
7.3312 | -34.2469 | -12.8525
7.3373 | -34.2472 | -12.8622
7.3434 | -34.2476 | -12.8722
7.3495 | -34.248 | -12.8819
7.3556 | -34.2483 | -12.8921
7.3617 | -34.2486 | -12.9023
7.3678 | -34.249 | -12.9121
7.3739 | -34.2493 | -12.9222
7.3801 | -34.2497 | -12.9326
7.3862 | -34.2501 | -12.9424
7.3923 | -34.2504 | -12.9524
7.3984 | -34.2508 | -12.9623
7.4045 | -34.2511 | -12.972
7.4106 | -34.2515 | -12.9819
7.4167 | -34.2518 | -12.9919
7.4228 | -34.2522 | -13.0018
7.4289 | -34.2526 | -13.0118
7.435 | -34.2529 | -13.0216
7.4411 | -34.2532 | -13.0315
7.4472 | -34.2536 | -13.0414
7.4533 | -34.254 | -13.051
7.4594 | -34.2543 | -13.0612
7.4655 | -34.2547 | -13.0709
7.4716 | -34.255 | -13.0811
7.4777 | -34.2554 | -13.0909
7.4838 | -34.2557 | -13.1005
7.4899 | -34.2561 | -13.11
7.496 | -34.2564 | -13.1197
7.5021 | -34.2568 | -13.1295
7.5082 | -34.2572 | -13.1394
7.5143 | -34.2575 | -13.1492
7.5204 | -34.2578 | -13.1592
7.5266 | -34.2582 | -13.169
7.5327 | -34.2586 | -13.1787
7.5388 | -34.2589 | -13.1889
7.5449 | -34.2593 | -13.199
7.551 | -34.2596 | -13.2087
7.5571 | -34.26 | -13.2189
7.5632 | -34.2604 | -13.2287
7.5693 | -34.2608 | -13.2387
7.5754 | -34.2611 | -13.2486
7.5815 | -34.2615 | -13.2588
7.5876 | -34.2618 | -13.2691
7.5937 | -34.2622 | -13.2792
7.5998 | -34.2626 | -13.289
7.6059 | -34.263 | -13.2989
7.612 | -34.2634 | -13.3088
7.6181 | -34.2637 | -13.3186
7.6242 | -34.2641 | -13.3288
7.6303 | -34.2645 | -13.339
7.6364 | -34.2648 | -13.349
7.6425 | -34.2652 | -13.3587
7.6486 | -34.2656 | -13.3687
7.6547 | -34.2659 | -13.3783
7.6608 | -34.2663 | -13.3883
7.667 | -34.2667 | -13.3983
7.6731 | -34.2671 | -13.4082
7.6792 | -34.2675 | -13.4184
7.6853 | -34.2679 | -13.4285
7.6914 | -34.2682 | -13.4386
7.6975 | -34.2686 | -13.4487
7.7036 | -34.2689 | -13.4586
7.7097 | -34.2693 | -13.4687
7.7158 | -34.2697 | -13.4786
7.7219 | -34.2701 | -13.4884
7.728 | -34.2705 | -13.4985
7.7341 | -34.2708 | -13.5083
7.7402 | -34.2712 | -13.5185
7.7463 | -34.2716 | -13.5284
7.7524 | -34.272 | -13.5387
7.7585 | -34.2723 | -13.5487
7.7646 | -34.2727 | -13.5586
7.7707 | -34.2731 | -13.5687
7.7768 | -34.2734 | -13.5788
7.7829 | -34.2738 | -13.589
7.789 | -34.2742 | -13.5992
7.7951 | -34.2745 | -13.6092
7.8012 | -34.2749 | -13.6191
7.8073 | -34.2753 | -13.6294
7.8135 | -34.2757 | -13.6391
7.8196 | -34.2761 | -13.6492
7.8257 | -34.2764 | -13.6594
7.8318 | -34.2768 | -13.6695
7.8379 | -34.2772 | -13.6795
7.844 | -34.2776 | -13.6894
7.8501 | -34.278 | -13.6996
7.8562 | -34.2783 | -13.7098
7.8623 | -34.2788 | -13.7198
7.8684 | -34.2791 | -13.7297
7.8745 | -34.2795 | -13.7398
7.8806 | -34.2799 | -13.7501
7.8867 | -34.2803 | -13.7599
7.8928 | -34.2806 | -13.7698
7.8989 | -34.281 | -13.7798
7.905 | -34.2814 | -13.7899
7.9111 | -34.2818 | -13.7998
7.9172 | -34.2822 | -13.8097
7.9233 | -34.2825 | -13.8199
7.9294 | -34.2829 | -13.83
7.9355 | -34.2833 | -13.8398
7.9416 | -34.2836 | -13.8498
7.9477 | -34.284 | -13.8597
7.9539 | -34.2844 | -13.8697
7.96 | -34.2847 | -13.8797
7.9661 | -34.2851 | -13.8897
7.9722 | -34.2855 | -13.8995
7.9783 | -34.2858 | -13.9099
7.9844 | -34.2863 | -13.9197
7.9905 | -34.2866 | -13.9299
7.9966 | -34.287 | -13.9401
8.0027 | -34.2873 | -13.9501
8.0088 | -34.2877 | -13.96
8.0149 | -34.288 | -13.9699
8.021 | -34.2884 | -13.9801
8.0271 | -34.2889 | -13.9904
8.0332 | -34.2892 | -14.0001
8.0393 | -34.2896 | -14.0103
8.0454 | -34.29 | -14.0205
8.0515 | -34.2905 | -14.0304
8.0576 | -34.2908 | -14.0406
8.0637 | -34.2912 | -14.0504
8.0698 | -34.2915 | -14.0605
8.0759 | -34.2919 | -14.0704
8.082 | -34.2923 | -14.0804
8.0881 | -34.2927 | -14.0904
8.0942 | -34.2931 | -14.1004
8.1004 | -34.2935 | -14.1104
8.1065 | -34.2938 | -14.1202
8.1126 | -34.2942 | -14.1302
8.1187 | -34.2946 | -14.1402
8.1248 | -34.295 | -14.1499
8.1309 | -34.2954 | -14.1597
8.137 | -34.2958 | -14.1695
8.1431 | -34.2961 | -14.1796
8.1492 | -34.2965 | -14.1898
8.1553 | -34.2969 | -14.1997
8.1614 | -34.2973 | -14.2099
8.1675 | -34.2976 | -14.2197
8.1736 | -34.298 | -14.23
8.1797 | -34.2984 | -14.2401
8.1858 | -34.2987 | -14.2502
8.1919 | -34.2991 | -14.2603
8.198 | -34.2995 | -14.2702
8.2041 | -34.2999 | -14.2802
8.2102 | -34.3002 | -14.2903
8.2163 | -34.3006 | -14.3002
8.2224 | -34.301 | -14.3102
8.2285 | -34.3014 | -14.3198
8.2346 | -34.3017 | -14.3298
8.2408 | -34.3021 | -14.3397
8.2469 | -34.3025 | -14.3497
8.253 | -34.303 | -14.3595
8.2591 | -34.3033 | -14.3696
8.2652 | -34.3038 | -14.3796
8.2713 | -34.3041 | -14.3894
8.2774 | -34.3045 | -14.3997
8.2835 | -34.3049 | -14.4098
8.2896 | -34.3053 | -14.4198
8.2957 | -34.3056 | -14.4298
8.3018 | -34.306 | -14.4397
8.3079 | -34.3064 | -14.4496
8.314 | -34.3068 | -14.4596
8.3201 | -34.3072 | -14.4694
8.3262 | -34.3075 | -14.4795
8.3323 | -34.3079 | -14.4895
8.3384 | -34.3083 | -14.4996
8.3445 | -34.3086 | -14.5096
8.3506 | -34.309 | -14.5196
8.3567 | -34.3094 | -14.5298
8.3628 | -34.3098 | -14.5398
8.3689 | -34.3102 | -14.5496
8.375 | -34.3106 | -14.5596
8.3812 | -34.3109 | -14.5693
8.3873 | -34.3114 | -14.5792
8.3934 | -34.3118 | -14.5892
8.3995 | -34.3121 | -14.5989
8.4056 | -34.3124 | -14.6087
8.4117 | -34.3128 | -14.6187
8.4178 | -34.3132 | -14.6287
8.4239 | -34.3135 | -14.6386
8.43 | -34.3139 | -14.6486
8.4361 | -34.3143 | -14.6585
8.4422 | -34.3147 | -14.6683
8.4483 | -34.3151 | -14.6781
8.4544 | -34.3155 | -14.6881
8.4605 | -34.3158 | -14.6977
8.4666 | -34.3162 | -14.708
8.4727 | -34.3166 | -14.7178
8.4788 | -34.317 | -14.7278
8.4849 | -34.3173 | -14.7378
8.491 | -34.3177 | -14.7479
8.4971 | -34.3181 | -14.7577
8.5032 | -34.3185 | -14.7676
8.5093 | -34.3189 | -14.7773
8.5154 | -34.3193 | -14.7873
8.5215 | -34.3197 | -14.7969
8.5277 | -34.32 | -14.8068
8.5338 | -34.3204 | -14.8167
8.5399 | -34.3208 | -14.8265
8.546 | -34.3212 | -14.8362
8.5521 | -34.3216 | -14.8461
8.5582 | -34.3219 | -14.856
8.5643 | -34.3223 | -14.8658
8.5704 | -34.3227 | -14.8755
8.5765 | -34.323 | -14.8853
8.5826 | -34.3234 | -14.8947
8.5887 | -34.3238 | -14.9046
8.5948 | -34.3242 | -14.9145
8.6009 | -34.3246 | -14.9242
8.607 | -34.325 | -14.9341
8.6131 | -34.3253 | -14.9439
8.6192 | -34.3258 | -14.9534
8.6253 | -34.3261 | -14.9632
8.6314 | -34.3265 | -14.973
8.6375 | -34.3269 | -14.9829
8.6436 | -34.3273 | -14.9926
8.6497 | -34.3277 | -15.0023
8.6558 | -34.328 | -15.0122
8.6619 | -34.3284 | -15.022
8.6681 | -34.3288 | -15.0321
8.6742 | -34.3292 | -15.0419
8.6803 | -34.3296 | -15.0516
8.6864 | -34.33 | -15.0612
8.6925 | -34.3304 | -15.0707
8.6986 | -34.3308 | -15.0805
8.7047 | -34.3311 | -15.0905
8.7108 | -34.3315 | -15.1002
8.7169 | -34.3319 | -15.11
8.723 | -34.3322 | -15.1195
8.7291 | -34.3326 | -15.1293
8.7352 | -34.333 | -15.1392
8.7413 | -34.3334 | -15.1491
8.7474 | -34.3337 | -15.1588
8.7535 | -34.3341 | -15.1686
8.7596 | -34.3345 | -15.1784
8.7657 | -34.3349 | -15.1882
8.7718 | -34.3353 | -15.198
8.7779 | -34.3357 | -15.2077
8.784 | -34.336 | -15.2177
8.7901 | -34.3365 | -15.2275
8.7962 | -34.3368 | -15.237
8.8023 | -34.3372 | -15.2467
8.8084 | -34.3376 | -15.2564
8.8146 | -34.3379 | -15.2663
8.8207 | -34.3383 | -15.2758
8.8268 | -34.3387 | -15.2852
8.8329 | -34.339 | -15.2949
8.839 | -34.3394 | -15.3046
8.8451 | -34.3398 | -15.3142
8.8512 | -34.3402 | -15.3237
8.8573 | -34.3405 | -15.3336
8.8634 | -34.3409 | -15.3431
8.8695 | -34.3413 | -15.3529
8.8756 | -34.3416 | -15.3628
8.8817 | -34.342 | -15.3721
8.8878 | -34.3424 | -15.3821
8.8939 | -34.3428 | -15.3918
8.9 | -34.3432 | -15.4014
8.9061 | -34.3435 | -15.4111
8.9122 | -34.3439 | -15.4208
8.9183 | -34.3443 | -15.4305
8.9244 | -34.3447 | -15.4404
8.9305 | -34.3451 | -15.4498
8.9366 | -34.3455 | -15.4596
8.9427 | -34.3459 | -15.4692
8.9488 | -34.3463 | -15.4789
8.955 | -34.3466 | -15.4885
8.9611 | -34.347 | -15.4983
8.9672 | -34.3474 | -15.5078
8.9733 | -34.3478 | -15.5174
8.9794 | -34.3482 | -15.5269
8.9855 | -34.3486 | -15.5366
8.9916 | -34.3489 | -15.5463
8.9977 | -34.3493 | -15.556
9.0038 | -34.3497 | -15.5655
9.0099 | -34.3501 | -15.5752
9.016 | -34.3505 | -15.5848
9.0221 | -34.3509 | -15.5944
9.0282 | -34.3513 | -15.6039
9.0343 | -34.3516 | -15.6134
9.0404 | -34.352 | -15.6234
9.0465 | -34.3524 | -15.6331
9.0526 | -34.3528 | -15.6427
9.0587 | -34.3532 | -15.6523
9.0648 | -34.3536 | -15.6617
9.0709 | -34.354 | -15.6712
9.077 | -34.3544 | -15.6807
9.0831 | -34.3548 | -15.6905
9.0892 | -34.3552 | -15.7003
9.0953 | -34.3556 | -15.7098
9.1015 | -34.3559 | -15.7195
9.1076 | -34.3563 | -15.7287
9.1137 | -34.3567 | -15.7383
9.1198 | -34.3571 | -15.7481
9.1259 | -34.3575 | -15.7576
9.132 | -34.3579 | -15.7675
9.1381 | -34.3583 | -15.777
9.1442 | -34.3588 | -15.7866
9.1503 | -34.3592 | -15.7965
9.1564 | -34.3595 | -15.8062
9.1625 | -34.3599 | -15.8159
9.1686 | -34.3603 | -15.8256
9.1747 | -34.3606 | -15.8354
9.1808 | -34.3611 | -15.8449
9.1869 | -34.3614 | -15.8545
9.193 | -34.3618 | -15.8641
9.1991 | -34.3621 | -15.8736
9.2052 | -34.3625 | -15.8832
9.2113 | -34.3629 | -15.8928
9.2174 | -34.3633 | -15.9024
9.2235 | -34.3637 | -15.9122
9.2296 | -34.3641 | -15.9216
9.2357 | -34.3644 | -15.9315
9.2419 | -34.3648 | -15.9412
9.248 | -34.3652 | -15.9505
9.2541 | -34.3656 | -15.9601
9.2602 | -34.366 | -15.9699
9.2663 | -34.3664 | -15.9793
9.2724 | -34.3668 | -15.9888
9.2785 | -34.3672 | -15.9984
9.2846 | -34.3676 | -16.0077
9.2907 | -34.368 | -16.0172
9.2968 | -34.3684 | -16.0268
9.3029 | -34.3688 | -16.0365
9.309 | -34.3692 | -16.0461
9.3151 | -34.3696 | -16.0554
9.3212 | -34.37 | -16.0649
9.3273 | -34.3704 | -16.0744
9.3334 | -34.3708 | -16.0837
9.3395 | -34.3712 | -16.0931
9.3456 | -34.3716 | -16.1026
9.3517 | -34.372 | -16.1122
9.3578 | -34.3723 | -16.1217
9.3639 | -34.3728 | -16.1313
9.37 | -34.3731 | -16.1409
9.3761 | -34.3735 | -16.1506
9.3822 | -34.3739 | -16.1602
9.3884 | -34.3743 | -16.1694
9.3945 | -34.3747 | -16.179
9.4006 | -34.3751 | -16.1883
9.4067 | -34.3755 | -16.1981
9.4128 | -34.3759 | -16.2077
9.4189 | -34.3763 | -16.2174
9.425 | -34.3767 | -16.2271
9.4311 | -34.3771 | -16.2364
9.4372 | -34.3775 | -16.246
9.4433 | -34.3779 | -16.2556
9.4494 | -34.3783 | -16.2648
9.4555 | -34.3787 | -16.2744
9.4616 | -34.379 | -16.2835
9.4677 | -34.3794 | -16.2927
9.4738 | -34.3798 | -16.3025
9.4799 | -34.3802 | -16.3119
9.486 | -34.3805 | -16.3213
9.4921 | -34.3809 | -16.3308
9.4982 | -34.3813 | -16.3402
9.5043 | -34.3817 | -16.3497
9.5104 | -34.3821 | -16.359
9.5165 | -34.3824 | -16.3686
9.5226 | -34.3828 | -16.3781
9.5288 | -34.3833 | -16.3875
9.5349 | -34.3836 | -16.3969
9.541 | -34.384 | -16.4061
9.5471 | -34.3844 | -16.4156
9.5532 | -34.3848 | -16.4251
9.5593 | -34.3852 | -16.4347
9.5654 | -34.3856 | -16.4439
9.5715 | -34.386 | -16.4533
9.5776 | -34.3864 | -16.4632
9.5837 | -34.3867 | -16.4726
9.5898 | -34.387 | -16.482
9.5959 | -34.3874 | -16.4918
9.602 | -34.3879 | -16.5011
9.6081 | -34.3883 | -16.5107
9.6142 | -34.3887 | -16.52
9.6203 | -34.389 | -16.5297
9.6264 | -34.3894 | -16.539
9.6325 | -34.3897 | -16.5488
9.6386 | -34.39 | -16.5583
9.6447 | -34.3905 | -16.5677
9.6508 | -34.3908 | -16.5769
9.6569 | -34.3912 | -16.5864
9.663 | -34.3916 | -16.596
9.6691 | -34.392 | -16.6056
9.6753 | -34.3924 | -16.6148
9.6814 | -34.3928 | -16.624
9.6875 | -34.3933 | -16.6332
9.6936 | -34.3937 | -16.6425
9.6997 | -34.3941 | -16.6517
9.7058 | -34.3945 | -16.6605
9.7119 | -34.395 | -16.6704
9.718 | -34.3954 | -16.6802
9.7241 | -34.3958 | -16.6896
9.7302 | -34.3961 | -16.6991
9.7363 | -34.3966 | -16.7085
9.7424 | -34.3969 | -16.7177
9.7485 | -34.3973 | -16.7274
9.7546 | -34.3977 | -16.7367
9.7607 | -34.3981 | -16.7465
9.7668 | -34.3985 | -16.756
9.7729 | -34.3989 | -16.7652
9.779 | -34.3994 | -16.774
9.7851 | -34.3997 | -16.7837
9.7912 | -34.4001 | -16.7932
9.7973 | -34.4005 | -16.8022
9.8034 | -34.4009 | -16.8114
9.8095 | -34.4013 | -16.8213
9.8157 | -34.4017 | -16.8313
9.8218 | -34.4021 | -16.8407
9.8279 | -34.4024 | -16.85
9.834 | -34.4027 | -16.8597
9.8401 | -34.4032 | -16.8692
9.8462 | -34.4036 | -16.8784
9.8523 | -34.4039 | -16.8878
9.8584 | -34.4043 | -16.8973
9.8645 | -34.4047 | -16.9069
9.8706 | -34.4051 | -16.9163
9.8767 | -34.4055 | -16.9257
9.8828 | -34.4058 | -16.9349
9.8889 | -34.4062 | -16.9441
9.895 | -34.4066 | -16.9536
9.9011 | -34.407 | -16.963
9.9072 | -34.4074 | -16.9725
9.9133 | -34.4078 | -16.9817
9.9194 | -34.4082 | -16.9911
9.9255 | -34.4085 | -17.0005
9.9316 | -34.4089 | -17.0099
9.9377 | -34.4093 | -17.0193
9.9438 | -34.4097 | -17.0286
9.9499 | -34.4101 | -17.0381
9.956 | -34.4105 | -17.0476
9.9622 | -34.4109 | -17.0573
9.9683 | -34.4113 | -17.0668
9.9744 | -34.4117 | -17.0763
9.9805 | -34.4121 | -17.0856
9.9866 | -34.4125 | -17.0949
9.9927 | -34.4128 | -17.1043
9.9988 | -34.4133 | -17.114
10.0049 | -34.4137 | -17.1232
10.011 | -34.4141 | -17.1328
10.0171 | -34.4145 | -17.1419
10.0232 | -34.4149 | -17.1511
10.0293 | -34.4153 | -17.1605
10.0354 | -34.4157 | -17.1699
10.0415 | -34.4161 | -17.1793
10.0476 | -34.4165 | -17.1888
10.0537 | -34.4169 | -17.1983
10.0598 | -34.4174 | -17.2077
10.0659 | -34.4177 | -17.2171
10.072 | -34.4181 | -17.2263
10.0781 | -34.4185 | -17.2355
10.0842 | -34.4189 | -17.245
10.0903 | -34.4193 | -17.2543
10.0964 | -34.4197 | -17.2637
10.1026 | -34.4201 | -17.2733
10.1087 | -34.4205 | -17.2827
10.1148 | -34.4208 | -17.292
10.1209 | -34.4213 | -17.3013
10.127 | -34.4217 | -17.3108
10.1331 | -34.4221 | -17.3198
10.1392 | -34.4225 | -17.3291
10.1453 | -34.4229 | -17.3383
10.1514 | -34.4233 | -17.3478
10.1575 | -34.4237 | -17.3573
10.1636 | -34.4241 | -17.3664
10.1697 | -34.4245 | -17.3757
10.1758 | -34.4249 | -17.3849
10.1819 | -34.4253 | -17.394
10.188 | -34.4256 | -17.4035
10.1941 | -34.426 | -17.4128
10.2002 | -34.4265 | -17.4223
10.2063 | -34.4269 | -17.4314
10.2124 | -34.4273 | -17.441
10.2185 | -34.4277 | -17.4505
10.2246 | -34.4281 | -17.4597
10.2307 | -34.4285 | -17.4693
10.2368 | -34.4289 | -17.4783
10.2429 | -34.4293 | -17.4876
10.2491 | -34.4297 | -17.4971
10.2552 | -34.4301 | -17.5064
10.2613 | -34.4305 | -17.5158
10.2674 | -34.431 | -17.5252
10.2735 | -34.4314 | -17.5345
10.2796 | -34.4318 | -17.5439
10.2857 | -34.4322 | -17.5536
10.2918 | -34.4326 | -17.5631
10.2979 | -34.433 | -17.5723
10.304 | -34.4335 | -17.5814
10.3101 | -34.4338 | -17.5908
10.3162 | -34.4343 | -17.6004
10.3223 | -34.4347 | -17.6095
10.3284 | -34.4351 | -17.6187
10.3345 | -34.4355 | -17.6279
10.3406 | -34.4359 | -17.6373
10.3467 | -34.4364 | -17.6466
10.3528 | -34.4368 | -17.6559
10.3589 | -34.4372 | -17.6652
10.365 | -34.4376 | -17.6748
10.3711 | -34.438 | -17.6842
10.3772 | -34.4385 | -17.6934
10.3833 | -34.4389 | -17.7026
10.3895 | -34.4394 | -17.7118
10.3956 | -34.4398 | -17.721
10.4017 | -34.4402 | -17.7303
10.4078 | -34.4406 | -17.7398
10.4139 | -34.441 | -17.7493
10.42 | -34.4415 | -17.7586
10.4261 | -34.4419 | -17.7678
10.4322 | -34.4422 | -17.7772
10.4383 | -34.4426 | -17.7865
10.4444 | -34.4431 | -17.7958
10.4505 | -34.4435 | -17.8051
10.4566 | -34.4439 | -17.8144
10.4627 | -34.4444 | -17.8234
10.4688 | -34.4448 | -17.8326
10.4749 | -34.4452 | -17.8419
10.481 | -34.4457 | -17.8511
10.4871 | -34.4461 | -17.8605
10.4932 | -34.4465 | -17.8696
10.4993 | -34.4469 | -17.8788
10.5054 | -34.4473 | -17.8882
10.5115 | -34.4478 | -17.8975
10.5176 | -34.4482 | -17.907
10.5237 | -34.4486 | -17.9163
10.5298 | -34.449 | -17.9255
10.536 | -34.4494 | -17.935
10.5421 | -34.4499 | -17.9442
10.5482 | -34.4503 | -17.9536
10.5543 | -34.4507 | -17.9627
10.5604 | -34.4512 | -17.9719
10.5665 | -34.4516 | -17.9809
10.5726 | -34.452 | -17.9903
10.5787 | -34.4524 | -17.9996
10.5848 | -34.4529 | -18.0085
10.5909 | -34.4533 | -18.018
10.597 | -34.4537 | -18.0274
10.6031 | -34.4542 | -18.0368
10.6092 | -34.4546 | -18.046
10.6153 | -34.455 | -18.0555
10.6214 | -34.4555 | -18.0648
10.6275 | -34.4559 | -18.0739
10.6336 | -34.4564 | -18.0833
10.6397 | -34.4568 | -18.0926
10.6458 | -34.4572 | -18.1019
10.6519 | -34.4576 | -18.1112
10.658 | -34.4581 | -18.1203
10.6641 | -34.4585 | -18.1296
10.6702 | -34.4589 | -18.1389
10.6764 | -34.4594 | -18.1482
10.6825 | -34.4599 | -18.1574
10.6886 | -34.4603 | -18.1664
10.6947 | -34.4608 | -18.1758
10.7008 | -34.4612 | -18.1851
10.7069 | -34.4616 | -18.1945
10.713 | -34.4621 | -18.2037
10.7191 | -34.4625 | -18.213
10.7252 | -34.463 | -18.222
10.7313 | -34.4634 | -18.2311
10.7374 | -34.4638 | -18.2406
10.7435 | -34.4643 | -18.2497
10.7496 | -34.4647 | -18.2588
10.7557 | -34.4651 | -18.2682
10.7618 | -34.4656 | -18.2773
10.7679 | -34.466 | -18.2864
10.774 | -34.4664 | -18.2956
10.7801 | -34.4669 | -18.3045
10.7862 | -34.4673 | -18.3135
10.7923 | -34.4678 | -18.3227
10.7984 | -34.4683 | -18.3319
10.8045 | -34.4687 | -18.3411
10.8106 | -34.4691 | -18.3505
10.8168 | -34.4695 | -18.3597
10.8229 | -34.47 | -18.3689
10.829 | -34.4704 | -18.378
10.8351 | -34.4709 | -18.3871
10.8412 | -34.4714 | -18.3962
10.8473 | -34.4718 | -18.4052
10.8534 | -34.4723 | -18.4143
10.8595 | -34.4727 | -18.4237
10.8656 | -34.4732 | -18.4327
10.8717 | -34.4736 | -18.442
10.8778 | -34.4741 | -18.4511
10.8839 | -34.4745 | -18.4603
10.89 | -34.4749 | -18.4695
10.8961 | -34.4754 | -18.4788
10.9022 | -34.4759 | -18.4878
10.9083 | -34.4763 | -18.497
10.9144 | -34.4768 | -18.506
10.9205 | -34.4772 | -18.5152
10.9266 | -34.4776 | -18.5244
10.9327 | -34.4781 | -18.5334
10.9388 | -34.4786 | -18.5424
10.9449 | -34.479 | -18.5516
10.951 | -34.4795 | -18.5609
10.9571 | -34.48 | -18.5701
10.9633 | -34.4804 | -18.5795
10.9694 | -34.4808 | -18.589
10.9755 | -34.4813 | -18.5981
10.9816 | -34.4817 | -18.6075
10.9877 | -34.4822 | -18.6163
10.9938 | -34.4827 | -18.6257
10.9999 | -34.4831 | -18.6349
11.006 | -34.4836 | -18.6438
11.0121 | -34.4841 | -18.6529
11.0182 | -34.4845 | -18.6621
11.0243 | -34.4849 | -18.6711
11.0304 | -34.4854 | -18.6803
11.0365 | -34.4858 | -18.6894
11.0426 | -34.4863 | -18.6985
11.0487 | -34.4868 | -18.7077
11.0548 | -34.4872 | -18.7168
11.0609 | -34.4877 | -18.726
11.067 | -34.4881 | -18.7354
11.0731 | -34.4886 | -18.7446
11.0792 | -34.489 | -18.7536
11.0853 | -34.4895 | -18.7626
11.0914 | -34.4899 | -18.7719
11.0975 | -34.4904 | -18.781
11.1037 | -34.4908 | -18.7901
11.1098 | -34.4913 | -18.7995
11.1159 | -34.4917 | -18.8084
11.122 | -34.4922 | -18.8176
11.1281 | -34.4926 | -18.8267
11.1342 | -34.4931 | -18.8358
11.1403 | -34.4935 | -18.8449
11.1464 | -34.494 | -18.8541
11.1525 | -34.4944 | -18.8631
11.1586 | -34.4949 | -18.8722
11.1647 | -34.4954 | -18.8812
11.1708 | -34.4958 | -18.8904
11.1769 | -34.4962 | -18.8994
11.183 | -34.4967 | -18.9085
11.1891 | -34.4972 | -18.9176
11.1952 | -34.4977 | -18.9266
11.2013 | -34.4981 | -18.9358
11.2074 | -34.4986 | -18.9449
11.2135 | -34.4991 | -18.9541
11.2196 | -34.4996 | -18.9632
11.2257 | -34.5 | -18.9722
11.2318 | -34.5005 | -18.9813
11.2379 | -34.501 | -18.9904
11.244 | -34.5015 | -18.9997
11.2502 | -34.5019 | -19.0088
11.2563 | -34.5024 | -19.0178
11.2624 | -34.5029 | -19.0271
11.2685 | -34.5033 | -19.0361
11.2746 | -34.5038 | -19.045
11.2807 | -34.5043 | -19.0542
11.2868 | -34.5047 | -19.0634
11.2929 | -34.5052 | -19.0721
11.299 | -34.5056 | -19.0814
11.3051 | -34.5061 | -19.0906
11.3112 | -34.5066 | -19.0997
11.3173 | -34.507 | -19.1086
11.3234 | -34.5075 | -19.1176
11.3295 | -34.508 | -19.1269
11.3356 | -34.5085 | -19.1363
11.3417 | -34.5089 | -19.1451
11.3478 | -34.5094 | -19.1544
11.3539 | -34.5098 | -19.1636
11.36 | -34.5103 | -19.1725
11.3661 | -34.5108 | -19.1816
11.3722 | -34.5112 | -19.1904
11.3783 | -34.5117 | -19.1995
11.3844 | -34.5122 | -19.2087
11.3906 | -34.5127 | -19.2179
11.3967 | -34.5131 | -19.2268
11.4028 | -34.5136 | -19.2359
11.4089 | -34.514 | -19.2448
11.415 | -34.5144 | -19.254
11.4211 | -34.5149 | -19.2633
11.4272 | -34.5153 | -19.2722
11.4333 | -34.5158 | -19.2813
11.4394 | -34.5162 | -19.2903
11.4455 | -34.5166 | -19.2993
11.4516 | -34.5171 | -19.3084
11.4577 | -34.5175 | -19.3176
11.4638 | -34.518 | -19.3264
11.4699 | -34.5184 | -19.3355
11.476 | -34.5189 | -19.3446
11.4821 | -34.5193 | -19.3535
11.4882 | -34.5198 | -19.3626
11.4943 | -34.5202 | -19.3717
11.5004 | -34.5207 | -19.3805
11.5065 | -34.5212 | -19.3895
11.5126 | -34.5216 | -19.3983
11.5187 | -34.5221 | -19.4074
11.5248 | -34.5225 | -19.4167
11.5309 | -34.523 | -19.4256
11.5371 | -34.5234 | -19.4347
11.5432 | -34.5239 | -19.4439
11.5493 | -34.5244 | -19.453
11.5554 | -34.5248 | -19.4621
11.5615 | -34.5253 | -19.4711
11.5676 | -34.5257 | -19.4801
11.5737 | -34.5261 | -19.489
11.5798 | -34.5266 | -19.4978
11.5859 | -34.5271 | -19.5066
11.592 | -34.5275 | -19.5157
11.5981 | -34.528 | -19.5245
11.6042 | -34.5284 | -19.5334
11.6103 | -34.5289 | -19.5423
11.6164 | -34.5293 | -19.5512
11.6225 | -34.5298 | -19.56
11.6286 | -34.5302 | -19.569
11.6347 | -34.5307 | -19.5781
11.6408 | -34.5311 | -19.587
11.6469 | -34.5316 | -19.5961
11.653 | -34.532 | -19.605
11.6591 | -34.5325 | -19.6138
11.6652 | -34.533 | -19.6227
11.6713 | -34.5334 | -19.6317
11.6775 | -34.5339 | -19.6403
11.6836 | -34.5343 | -19.6492
11.6897 | -34.5348 | -19.6582
11.6958 | -34.5352 | -19.6672
11.7019 | -34.5356 | -19.6762
11.708 | -34.5361 | -19.6852
11.7141 | -34.5366 | -19.6939
11.7202 | -34.537 | -19.7029
11.7263 | -34.5374 | -19.7119
11.7324 | -34.5379 | -19.721
11.7385 | -34.5383 | -19.7298
11.7446 | -34.5387 | -19.7386
11.7507 | -34.5392 | -19.7474
11.7568 | -34.5396 | -19.7564
11.7629 | -34.5401 | -19.7656
11.769 | -34.5405 | -19.7743
11.7751 | -34.541 | -19.7834
11.7812 | -34.5414 | -19.7921
11.7873 | -34.5418 | -19.8009
11.7934 | -34.5423 | -19.8099
11.7995 | -34.5428 | -19.8188
11.8056 | -34.5432 | -19.8277
11.8117 | -34.5437 | -19.8367
11.8178 | -34.5441 | -19.8454
11.824 | -34.5446 | -19.8543
11.8301 | -34.545 | -19.8632
11.8362 | -34.5454 | -19.872
11.8423 | -34.5459 | -19.8808
11.8484 | -34.5463 | -19.8897
11.8545 | -34.5467 | -19.8985
11.8606 | -34.5472 | -19.9073
11.8667 | -34.5476 | -19.9161
11.8728 | -34.5481 | -19.9249
11.8789 | -34.5485 | -19.9338
11.885 | -34.549 | -19.9426
11.8911 | -34.5494 | -19.9514
11.8972 | -34.5498 | -19.9603
11.9033 | -34.5503 | -19.9692
11.9094 | -34.5507 | -19.9781
11.9155 | -34.5512 | -19.9869
11.9216 | -34.5516 | -19.9955
11.9277 | -34.552 | -20.0043
11.9338 | -34.5525 | -20.0134
11.9399 | -34.5529 | -20.0225
11.946 | -34.5534 | -20.0313
11.9521 | -34.5538 | -20.0401
11.9582 | -34.5543 | -20.0488
11.9644 | -34.5547 | -20.0575
11.9705 | -34.5552 | -20.0664
11.9766 | -34.5556 | -20.0752
11.9827 | -34.5561 | -20.0841
11.9888 | -34.5566 | -20.0929
11.9949 | -34.557 | -20.1012
12.001 | -34.5574 | -20.1101
12.0071 | -34.5579 | -20.1187
12.0132 | -34.5583 | -20.1274
12.0193 | -34.5587 | -20.1364
12.0254 | -34.5591 | -20.1452
12.0315 | -34.5596 | -20.1538
12.0376 | -34.56 | -20.1625
12.0437 | -34.5604 | -20.1714
12.0498 | -34.5609 | -20.1802
12.0559 | -34.5613 | -20.1888
12.062 | -34.5618 | -20.1975
12.0681 | -34.5622 | -20.2064
12.0742 | -34.5627 | -20.215
12.0803 | -34.5631 | -20.2238
12.0864 | -34.5635 | -20.2327
12.0925 | -34.564 | -20.2415
12.0986 | -34.5644 | -20.25
12.1047 | -34.5648 | -20.2588
12.1109 | -34.5653 | -20.2675
12.117 | -34.5657 | -20.2763
12.1231 | -34.5662 | -20.285
12.1292 | -34.5666 | -20.2937
12.1353 | -34.567 | -20.3024
12.1414 | -34.5674 | -20.3112
12.1475 | -34.5679 | -20.32
12.1536 | -34.5683 | -20.3289
12.1597 | -34.5687 | -20.3376
12.1658 | -34.5692 | -20.3464
12.1719 | -34.5696 | -20.3548
12.178 | -34.57 | -20.3635
12.1841 | -34.5705 | -20.3722
12.1902 | -34.5709 | -20.3811
12.1963 | -34.5714 | -20.39
12.2024 | -34.5718 | -20.3985
12.2085 | -34.5722 | -20.4073
12.2146 | -34.5727 | -20.416
12.2207 | -34.5731 | -20.4249
12.2268 | -34.5735 | -20.4337
12.2329 | -34.574 | -20.4424
12.239 | -34.5744 | -20.4508
12.2451 | -34.5749 | -20.4596
12.2513 | -34.5753 | -20.4685
12.2574 | -34.5758 | -20.4773
12.2635 | -34.5762 | -20.486
12.2696 | -34.5766 | -20.4946
12.2757 | -34.5771 | -20.5031
12.2818 | -34.5775 | -20.5117
12.2879 | -34.5779 | -20.5205
12.294 | -34.5784 | -20.5291
12.3001 | -34.5788 | -20.5378
12.3062 | -34.5793 | -20.5465
12.3123 | -34.5797 | -20.5549
12.3184 | -34.5801 | -20.5635
12.3245 | -34.5806 | -20.5721
12.3306 | -34.581 | -20.5808
12.3367 | -34.5814 | -20.5896
12.3428 | -34.5818 | -20.5983
12.3489 | -34.5823 | -20.6068
12.355 | -34.5827 | -20.6157
12.3611 | -34.5831 | -20.6245
12.3672 | -34.5835 | -20.6332
12.3733 | -34.584 | -20.6419
12.3794 | -34.5844 | -20.6504
12.3855 | -34.5849 | -20.659
12.3916 | -34.5853 | -20.6676
12.3978 | -34.5857 | -20.6764
12.4039 | -34.5862 | -20.6851
12.41 | -34.5866 | -20.6938
12.4161 | -34.5871 | -20.7021
12.4222 | -34.5875 | -20.7109
12.4283 | -34.5879 | -20.7195
12.4344 | -34.5884 | -20.7282
12.4405 | -34.5888 | -20.737
12.4466 | -34.5892 | -20.7455
12.4527 | -34.5897 | -20.754
12.4588 | -34.5901 | -20.7627
12.4649 | -34.5905 | -20.7714
12.471 | -34.591 | -20.7801
12.4771 | -34.5914 | -20.7887
12.4832 | -34.5918 | -20.7972
12.4893 | -34.5922 | -20.8057
12.4954 | -34.5926 | -20.8141
12.5015 | -34.5931 | -20.8225
12.5076 | -34.5935 | -20.8312
12.5137 | -34.594 | -20.8398
12.5198 | -34.5944 | -20.8484
12.5259 | -34.5948 | -20.8571
12.532 | -34.5952 | -20.8657
12.5382 | -34.5957 | -20.8745
12.5443 | -34.5961 | -20.8831
12.5504 | -34.5965 | -20.8918
12.5565 | -34.597 | -20.9004
12.5626 | -34.5974 | -20.9092
12.5687 | -34.5979 | -20.9178
12.5748 | -34.5983 | -20.9266
12.5809 | -34.5987 | -20.9352
12.587 | -34.5991 | -20.9439
12.5931 | -34.5996 | -20.9522
12.5992 | -34.6001 | -20.961
12.6053 | -34.6005 | -20.9695
12.6114 | -34.6009 | -20.9781
12.6175 | -34.6013 | -20.9867
12.6236 | -34.6018 | -20.9956
12.6297 | -34.6022 | -21.004
12.6358 | -34.6026 | -21.0124
12.6419 | -34.6031 | -21.021
12.648 | -34.6035 | -21.0295
12.6541 | -34.604 | -21.0382
12.6602 | -34.6044 | -21.0465
12.6663 | -34.6048 | -21.0551
12.6724 | -34.6053 | -21.0637
12.6785 | -34.6057 | -21.0723
12.6847 | -34.6061 | -21.0807
12.6908 | -34.6066 | -21.0894
12.6969 | -34.607 | -21.0979
12.703 | -34.6074 | -21.1064
12.7091 | -34.6079 | -21.1149
12.7152 | -34.6083 | -21.1235
12.7213 | -34.6088 | -21.132
12.7274 | -34.6092 | -21.1406
12.7335 | -34.6096 | -21.149
12.7396 | -34.6101 | -21.1574
12.7457 | -34.6105 | -21.166
12.7518 | -34.6109 | -21.1746
12.7579 | -34.6113 | -21.1833
12.764 | -34.6118 | -21.192
12.7701 | -34.6122 | -21.2005
12.7762 | -34.6126 | -21.2091
12.7823 | -34.6131 | -21.2177
12.7884 | -34.6135 | -21.2262
12.7945 | -34.614 | -21.2351
12.8006 | -34.6144 | -21.2435
12.8067 | -34.6149 | -21.252
12.8128 | -34.6153 | -21.2603
12.8189 | -34.6158 | -21.269
12.8251 | -34.6162 | -21.2775
12.8312 | -34.6166 | -21.2861
12.8373 | -34.6171 | -21.2946
12.8434 | -34.6175 | -21.3033
12.8495 | -34.618 | -21.3117
12.8556 | -34.6184 | -21.3201
12.8617 | -34.6188 | -21.3287
12.8678 | -34.6193 | -21.3375
12.8739 | -34.6197 | -21.3459
12.88 | -34.6201 | -21.3544
12.8861 | -34.6205 | -21.363
12.8922 | -34.621 | -21.3717
12.8983 | -34.6214 | -21.3804
12.9044 | -34.6218 | -21.3888
12.9105 | -34.6223 | -21.3974
12.9166 | -34.6227 | -21.4059
12.9227 | -34.6232 | -21.4144
12.9288 | -34.6236 | -21.423
12.9349 | -34.6241 | -21.4316
12.941 | -34.6246 | -21.4401
12.9471 | -34.625 | -21.4488
12.9532 | -34.6255 | -21.4573
12.9593 | -34.6259 | -21.466
12.9654 | -34.6264 | -21.4745
12.9716 | -34.6268 | -21.4832
12.9777 | -34.6273 | -21.4917
12.9838 | -34.6278 | -21.5001
12.9899 | -34.6281 | -21.5084
12.996 | -34.6286 | -21.5168
13.0021 | -34.629 | -21.5253
13.0082 | -34.6295 | -21.5336
13.0143 | -34.63 | -21.5419
13.0204 | -34.6304 | -21.5502
13.0265 | -34.6309 | -21.5586
13.0326 | -34.6313 | -21.5672
13.0387 | -34.6318 | -21.5757
13.0448 | -34.6322 | -21.5845
13.0509 | -34.6327 | -21.5928
13.057 | -34.6331 | -21.6011
13.0631 | -34.6336 | -21.6096
13.0692 | -34.6341 | -21.6182
13.0753 | -34.6345 | -21.6269
13.0814 | -34.635 | -21.6356
13.0875 | -34.6354 | -21.6439
13.0936 | -34.6358 | -21.6526
13.0997 | -34.6363 | -21.6609
13.1058 | -34.6368 | -21.6694
13.112 | -34.6372 | -21.6779
13.1181 | -34.6377 | -21.6865
13.1242 | -34.6381 | -21.6951
13.1303 | -34.6386 | -21.7035
13.1364 | -34.639 | -21.712
13.1425 | -34.6395 | -21.7205
13.1486 | -34.6399 | -21.7292
13.1547 | -34.6404 | -21.7376
13.1608 | -34.6408 | -21.746
13.1669 | -34.6413 | -21.7544
13.173 | -34.6417 | -21.763
13.1791 | -34.6422 | -21.7715
13.1852 | -34.6426 | -21.78
13.1913 | -34.6431 | -21.7886
13.1974 | -34.6435 | -21.7973
13.2035 | -34.644 | -21.8057
13.2096 | -34.6444 | -21.8143
13.2157 | -34.6449 | -21.8227
13.2218 | -34.6453 | -21.831
13.2279 | -34.6458 | -21.8394
13.234 | -34.6463 | -21.8478
13.2401 | -34.6467 | -21.8565
13.2462 | -34.6471 | -21.8649
13.2524 | -34.6476 | -21.8735
13.2585 | -34.6481 | -21.8819
13.2646 | -34.6485 | -21.8904
13.2707 | -34.649 | -21.8989
13.2768 | -34.6494 | -21.9076
13.2829 | -34.6499 | -21.9161
13.289 | -34.6503 | -21.9246
13.2951 | -34.6508 | -21.933
13.3012 | -34.6512 | -21.9416
13.3073 | -34.6517 | -21.9502
13.3134 | -34.6522 | -21.9586
13.3195 | -34.6526 | -21.967
13.3256 | -34.6531 | -21.9755
13.3317 | -34.6536 | -21.984
13.3378 | -34.654 | -21.9924
13.3439 | -34.6545 | -22.0011
13.35 | -34.6549 | -22.0096
13.3561 | -34.6554 | -22.0181
13.3622 | -34.6558 | -22.0266
13.3683 | -34.6563 | -22.0349
13.3744 | -34.6568 | -22.0431
13.3805 | -34.6572 | -22.0516
13.3866 | -34.6577 | -22.0599
13.3927 | -34.6582 | -22.0683
13.3989 | -34.6586 | -22.0766
13.405 | -34.6591 | -22.085
13.4111 | -34.6595 | -22.0934
13.4172 | -34.6599 | -22.1019
13.4233 | -34.6604 | -22.1105
13.4294 | -34.6609 | -22.1193
13.4355 | -34.6613 | -22.1276
13.4416 | -34.6618 | -22.136
13.4477 | -34.6623 | -22.1446
13.4538 | -34.6628 | -22.1532
13.4599 | -34.6632 | -22.1619
13.466 | -34.6637 | -22.1703
13.4721 | -34.6642 | -22.1787
13.4782 | -34.6646 | -22.1872
13.4843 | -34.6651 | -22.1959
13.4904 | -34.6656 | -22.2043
13.4965 | -34.666 | -22.2126
13.5026 | -34.6665 | -22.2212
13.5087 | -34.667 | -22.2296
13.5148 | -34.6674 | -22.2381
13.5209 | -34.6679 | -22.2467
13.527 | -34.6683 | -22.2549
13.5331 | -34.6688 | -22.2635
13.5393 | -34.6693 | -22.2718
13.5454 | -34.6697 | -22.2797
13.5515 | -34.6702 | -22.2886
13.5576 | -34.6707 | -22.2972
13.5637 | -34.6711 | -22.3057
13.5698 | -34.6716 | -22.3142
13.5759 | -34.6721 | -22.3224
13.582 | -34.6725 | -22.3308
13.5881 | -34.673 | -22.3393
13.5942 | -34.6734 | -22.3478
13.6003 | -34.6739 | -22.3563
13.6064 | -34.6744 | -22.365
13.6125 | -34.6749 | -22.3734
13.6186 | -34.6754 | -22.3816
13.6247 | -34.6759 | -22.3901
13.6308 | -34.6764 | -22.3988
13.6369 | -34.6768 | -22.4071
13.643 | -34.6773 | -22.4155
13.6491 | -34.6777 | -22.4238
13.6552 | -34.6782 | -22.4323
13.6613 | -34.6786 | -22.4406
13.6674 | -34.6791 | -22.4492
13.6735 | -34.6796 | -22.4574
13.6796 | -34.6801 | -22.4658
13.6858 | -34.6805 | -22.4741
13.6919 | -34.681 | -22.4824
13.698 | -34.6816 | -22.491
13.7041 | -34.682 | -22.4993
13.7102 | -34.6825 | -22.5078
13.7163 | -34.683 | -22.5161
13.7224 | -34.6834 | -22.5247
13.7285 | -34.6839 | -22.5331
13.7346 | -34.6844 | -22.5415
13.7407 | -34.6849 | -22.5498
13.7468 | -34.6853 | -22.5583
13.7529 | -34.6858 | -22.5666
13.759 | -34.6863 | -22.5752
13.7651 | -34.6867 | -22.5837
13.7712 | -34.6871 | -22.5922
13.7773 | -34.6875 | -22.6006
13.7834 | -34.688 | -22.6092
13.7895 | -34.6885 | -22.6175
13.7956 | -34.689 | -22.6262
13.8017 | -34.6894 | -22.6347
13.8078 | -34.6899 | -22.6427
13.8139 | -34.6903 | -22.6508
13.82 | -34.6908 | -22.6592
13.8262 | -34.6913 | -22.6676
13.8323 | -34.6918 | -22.6761
13.8384 | -34.6923 | -22.6846
13.8445 | -34.6928 | -22.6931
13.8506 | -34.6933 | -22.7014
13.8567 | -34.6937 | -22.7098
13.8628 | -34.6942 | -22.7181
13.8689 | -34.6947 | -22.7266
13.875 | -34.6952 | -22.7349
13.8811 | -34.6956 | -22.7432
13.8872 | -34.6961 | -22.7516
13.8933 | -34.6966 | -22.7598
13.8994 | -34.697 | -22.7683
13.9055 | -34.6975 | -22.7772
13.9116 | -34.698 | -22.7854
13.9177 | -34.6985 | -22.7936
13.9238 | -34.699 | -22.8024
13.9299 | -34.6995 | -22.8107
13.936 | -34.7 | -22.8191
13.9421 | -34.7004 | -22.8272
13.9482 | -34.7009 | -22.8353
13.9543 | -34.7014 | -22.8435
13.9604 | -34.7019 | -22.852
13.9665 | -34.7024 | -22.8603
13.9727 | -34.7029 | -22.8688
13.9788 | -34.7034 | -22.8769
13.9849 | -34.7038 | -22.8852
13.991 | -34.7043 | -22.8935
13.9971 | -34.7048 | -22.9016
14.0032 | -34.7053 | -22.91
14.0093 | -34.7058 | -22.9185
14.0154 | -34.7062 | -22.9268
14.0215 | -34.7067 | -22.9351
14.0276 | -34.7072 | -22.9435
14.0337 | -34.7077 | -22.9518
14.0398 | -34.7082 | -22.9602
14.0459 | -34.7086 | -22.9686
14.052 | -34.7091 | -22.9771
14.0581 | -34.7096 | -22.9856
14.0642 | -34.7101 | -22.994
14.0703 | -34.7106 | -23.0023
14.0764 | -34.7111 | -23.0104
14.0825 | -34.7116 | -23.0187
14.0886 | -34.7121 | -23.0274
14.0947 | -34.7126 | -23.0355
14.1008 | -34.713 | -23.0439
14.1069 | -34.7135 | -23.052
14.1131 | -34.714 | -23.0605
14.1192 | -34.7145 | -23.0687
14.1253 | -34.715 | -23.0769
14.1314 | -34.7154 | -23.0853
14.1375 | -34.7159 | -23.0935
14.1436 | -34.7164 | -23.1018
14.1497 | -34.7169 | -23.11
14.1558 | -34.7174 | -23.1184
14.1619 | -34.7179 | -23.1266
14.168 | -34.7184 | -23.135
14.1741 | -34.7189 | -23.1431
14.1802 | -34.7194 | -23.1514
14.1863 | -34.7198 | -23.1599
14.1924 | -34.7203 | -23.1682
14.1985 | -34.7209 | -23.1765
14.2046 | -34.7213 | -23.185
14.2107 | -34.7219 | -23.193
14.2168 | -34.7223 | -23.2012
14.2229 | -34.7228 | -23.2097
14.229 | -34.7233 | -23.218
14.2351 | -34.7238 | -23.2264
14.2412 | -34.7243 | -23.2349
14.2473 | -34.7248 | -23.243
14.2534 | -34.7253 | -23.2511
14.2596 | -34.7258 | -23.2594
14.2657 | -34.7262 | -23.2676
14.2718 | -34.7267 | -23.2759
14.2779 | -34.7272 | -23.2842
14.284 | -34.7277 | -23.2925
14.2901 | -34.7282 | -23.3009
14.2962 | -34.7287 | -23.3091
14.3023 | -34.7291 | -23.3174
14.3084 | -34.7296 | -23.3258
14.3145 | -34.7302 | -23.3339
14.3206 | -34.7307 | -23.3423
14.3267 | -34.7312 | -23.3507
14.3328 | -34.7317 | -23.359
14.3389 | -34.7321 | -23.3672
14.345 | -34.7326 | -23.3755
14.3511 | -34.7331 | -23.3839
14.3572 | -34.7336 | -23.3921
14.3633 | -34.7341 | -23.4003
14.3694 | -34.7346 | -23.4085
14.3755 | -34.7351 | -23.4169
14.3816 | -34.7356 | -23.4252
14.3877 | -34.7361 | -23.4332
14.3938 | -34.7366 | -23.4414
14.4 | -34.7371 | -23.4496
14.4061 | -34.7376 | -23.458
14.4122 | -34.738 | -23.4664
14.4183 | -34.7385 | -23.4746
14.4244 | -34.739 | -23.4828
14.4305 | -34.7395 | -23.491
14.4366 | -34.7401 | -23.4992
14.4427 | -34.7406 | -23.5078
14.4488 | -34.7411 | -23.516
14.4549 | -34.7416 | -23.5243
14.461 | -34.7421 | -23.5325
14.4671 | -34.7426 | -23.541
14.4732 | -34.743 | -23.5493
14.4793 | -34.7435 | -23.5577
14.4854 | -34.7441 | -23.566
14.4915 | -34.7446 | -23.5739
14.4976 | -34.7451 | -23.5819
14.5037 | -34.7456 | -23.59
14.5098 | -34.7461 | -23.5984
14.5159 | -34.7466 | -23.6066
14.522 | -34.747 | -23.6147
14.5281 | -34.7475 | -23.6227
14.5342 | -34.7481 | -23.6311
14.5403 | -34.7486 | -23.6392
14.5465 | -34.7491 | -23.6476
14.5526 | -34.7495 | -23.656
14.5587 | -34.75 | -23.6641
14.5648 | -34.7505 | -23.6721
14.5709 | -34.751 | -23.6804
14.577 | -34.7515 | -23.6886
14.5831 | -34.752 | -23.6968
14.5892 | -34.7525 | -23.705
14.5953 | -34.753 | -23.713
14.6014 | -34.7535 | -23.7213
14.6075 | -34.754 | -23.7295
14.6136 | -34.7544 | -23.7377
14.6197 | -34.7549 | -23.7459
14.6258 | -34.7554 | -23.7542
14.6319 | -34.7559 | -23.7623
14.638 | -34.7564 | -23.7705
14.6441 | -34.7569 | -23.7787
14.6502 | -34.7574 | -23.7871
14.6563 | -34.7579 | -23.7954
14.6624 | -34.7584 | -23.8035
14.6685 | -34.7589 | -23.8115
14.6746 | -34.7594 | -23.8198
14.6807 | -34.7599 | -23.828
14.6869 | -34.7604 | -23.8362
14.693 | -34.7609 | -23.8442
14.6991 | -34.7614 | -23.8524
14.7052 | -34.762 | -23.8604
14.7113 | -34.7625 | -23.8686
14.7174 | -34.763 | -23.8768
14.7235 | -34.7635 | -23.8849
14.7296 | -34.7639 | -23.893
14.7357 | -34.7645 | -23.9011
14.7418 | -34.765 | -23.9094
14.7479 | -34.7654 | -23.9174
14.754 | -34.7659 | -23.9258
14.7601 | -34.7664 | -23.9339
14.7662 | -34.7669 | -23.9421
14.7723 | -34.7674 | -23.9499
14.7784 | -34.7679 | -23.9581
14.7845 | -34.7684 | -23.9663
14.7906 | -34.7689 | -23.9744
14.7967 | -34.7694 | -23.9826
14.8028 | -34.7699 | -23.9906
14.8089 | -34.7704 | -23.9987
14.815 | -34.7709 | -24.0068
14.8211 | -34.7714 | -24.015
14.8272 | -34.7719 | -24.0233
14.8334 | -34.7724 | -24.0314
14.8395 | -34.7729 | -24.0395
14.8456 | -34.7734 | -24.0475
14.8517 | -34.7739 | -24.0556
14.8578 | -34.7744 | -24.0639
14.8639 | -34.7749 | -24.0721
14.87 | -34.7754 | -24.0804
14.8761 | -34.7759 | -24.0886
14.8822 | -34.7764 | -24.0968
14.8883 | -34.7769 | -24.1051
14.8944 | -34.7774 | -24.1132
14.9005 | -34.7779 | -24.1215
14.9066 | -34.7784 | -24.1296
14.9127 | -34.7789 | -24.1376
14.9188 | -34.7794 | -24.1457
14.9249 | -34.7799 | -24.1539
14.931 | -34.7804 | -24.1621
14.9371 | -34.7809 | -24.1701
14.9432 | -34.7814 | -24.1783
14.9493 | -34.7819 | -24.1864
14.9554 | -34.7824 | -24.1945
14.9615 | -34.7829 | -24.2027
14.9676 | -34.7834 | -24.2111
14.9738 | -34.7839 | -24.219
14.9799 | -34.7844 | -24.2273
14.986 | -34.7849 | -24.2353
14.9921 | -34.7854 | -24.2434
14.9982 | -34.7859 | -24.2515
15.0043 | -34.7864 | -24.2596
15.0104 | -34.7869 | -24.268
15.0165 | -34.7874 | -24.2765
15.0226 | -34.7879 | -24.2844
15.0287 | -34.7884 | -24.2926
15.0348 | -34.7889 | -24.3008
15.0409 | -34.7894 | -24.3091
15.047 | -34.7899 | -24.3173
15.0531 | -34.7905 | -24.3253
15.0592 | -34.7909 | -24.3335
15.0653 | -34.7915 | -24.3415
15.0714 | -34.7919 | -24.3497
15.0775 | -34.7924 | -24.3577
15.0836 | -34.793 | -24.3658
15.0897 | -34.7935 | -24.3737
15.0958 | -34.794 | -24.3817
15.1019 | -34.7945 | -24.3898
15.108 | -34.795 | -24.398
15.1141 | -34.7955 | -24.406
15.1203 | -34.796 | -24.4142
15.1264 | -34.7966 | -24.4222
15.1325 | -34.797 | -24.4303
15.1386 | -34.7975 | -24.4384
15.1447 | -34.7981 | -24.4466
15.1508 | -34.7986 | -24.4548
15.1569 | -34.7991 | -24.463
15.163 | -34.7996 | -24.4708
15.1691 | -34.8001 | -24.4788
15.1752 | -34.8006 | -24.4868
15.1813 | -34.8011 | -24.4947
15.1874 | -34.8016 | -24.5029
15.1935 | -34.8021 | -24.5108
15.1996 | -34.8026 | -24.5188
15.2057 | -34.8031 | -24.5268
15.2118 | -34.8036 | -24.5349
15.2179 | -34.8041 | -24.543
15.224 | -34.8046 | -24.5511
15.2301 | -34.8051 | -24.5588
15.2362 | -34.8056 | -24.5668
15.2423 | -34.8061 | -24.5751
15.2484 | -34.8066 | -24.5832
15.2545 | -34.8071 | -24.5914
15.2607 | -34.8076 | -24.5994
15.2668 | -34.8081 | -24.6073
15.2729 | -34.8086 | -24.6154
15.279 | -34.8092 | -24.6236
15.2851 | -34.8097 | -24.6316
15.2912 | -34.8102 | -24.6397
15.2973 | -34.8107 | -24.6478
15.3034 | -34.8112 | -24.6558
15.3095 | -34.8117 | -24.6639
15.3156 | -34.8122 | -24.6719
15.3217 | -34.8127 | -24.68
15.3278 | -34.8132 | -24.6882
15.3339 | -34.8138 | -24.6961
15.34 | -34.8143 | -24.704
15.3461 | -34.8148 | -24.7119
15.3522 | -34.8152 | -24.7198
15.3583 | -34.8157 | -24.728
15.3644 | -34.8162 | -24.7359
15.3705 | -34.8167 | -24.7437
15.3766 | -34.8172 | -24.7519
15.3827 | -34.8178 | -24.7599
15.3888 | -34.8183 | -24.768
15.3949 | -34.8188 | -24.7761
15.401 | -34.8193 | -24.7842
15.4072 | -34.8198 | -24.792
15.4133 | -34.8202 | -24.8001
15.4194 | -34.8207 | -24.8081
15.4255 | -34.8212 | -24.8163
15.4316 | -34.8218 | -24.8243
15.4377 | -34.8223 | -24.8324
15.4438 | -34.8228 | -24.8404
15.4499 | -34.8233 | -24.8481
15.456 | -34.8238 | -24.856
15.4621 | -34.8244 | -24.864
15.4682 | -34.8248 | -24.8722
15.4743 | -34.8253 | -24.8802
15.4804 | -34.8258 | -24.8882
15.4865 | -34.8263 | -24.8962
15.4926 | -34.8268 | -24.9043
15.4987 | -34.8273 | -24.9123
15.5048 | -34.8278 | -24.9203
15.5109 | -34.8283 | -24.9282
15.517 | -34.8288 | -24.9362
15.5231 | -34.8293 | -24.9442
15.5292 | -34.8298 | -24.9521
15.5353 | -34.8303 | -24.9602
15.5414 | -34.8308 | -24.968
15.5476 | -34.8314 | -24.9757
15.5537 | -34.8319 | -24.9837
15.5598 | -34.8324 | -24.9916
15.5659 | -34.8329 | -24.9996
15.572 | -34.8334 | -25.0075
15.5781 | -34.8339 | -25.0154
15.5842 | -34.8344 | -25.0232
15.5903 | -34.8349 | -25.0311
15.5964 | -34.8354 | -25.039
15.6025 | -34.8359 | -25.0468
15.6086 | -34.8364 | -25.0548
15.6147 | -34.8369 | -25.063
15.6208 | -34.8374 | -25.0707
15.6269 | -34.8379 | -25.0785
15.633 | -34.8384 | -25.0865
15.6391 | -34.8389 | -25.0946
15.6452 | -34.8394 | -25.1027
15.6513 | -34.84 | -25.1103
15.6574 | -34.8405 | -25.1185
15.6635 | -34.841 | -25.1265
15.6696 | -34.8415 | -25.1345
15.6757 | -34.842 | -25.1425
15.6818 | -34.8425 | -25.1506
15.688 | -34.843 | -25.1582
15.6941 | -34.8436 | -25.1661
15.7002 | -34.8441 | -25.174
15.7063 | -34.8446 | -25.1819
15.7124 | -34.8451 | -25.1899
15.7185 | -34.8456 | -25.1978
15.7246 | -34.8461 | -25.2057
15.7307 | -34.8466 | -25.2135
15.7368 | -34.8472 | -25.2214
15.7429 | -34.8477 | -25.2292
15.749 | -34.8482 | -25.2372
15.7551 | -34.8488 | -25.2452
15.7612 | -34.8493 | -25.253
15.7673 | -34.8498 | -25.2607
15.7734 | -34.8503 | -25.2686
15.7795 | -34.8508 | -25.2764
15.7856 | -34.8512 | -25.2842
15.7917 | -34.8518 | -25.2918
15.7978 | -34.8523 | -25.2999
15.8039 | -34.8528 | -25.3079
15.81 | -34.8533 | -25.3157
15.8161 | -34.8538 | -25.3236
15.8222 | -34.8543 | -25.3315
15.8283 | -34.8549 | -25.3391
15.8345 | -34.8554 | -25.3472
15.8406 | -34.8559 | -25.3551
15.8467 | -34.8564 | -25.363
15.8528 | -34.8569 | -25.371
15.8589 | -34.8574 | -25.3789
15.865 | -34.858 | -25.3864
15.8711 | -34.8585 | -25.3944
15.8772 | -34.859 | -25.4021
15.8833 | -34.8595 | -25.4099
15.8894 | -34.86 | -25.4179
15.8955 | -34.8605 | -25.4257
15.9016 | -34.861 | -25.4332
15.9077 | -34.8615 | -25.4412
15.9138 | -34.862 | -25.4491
15.9199 | -34.8625 | -25.4569
15.926 | -34.863 | -25.4646
15.9321 | -34.8635 | -25.4726
15.9382 | -34.864 | -25.4804
15.9443 | -34.8645 | -25.4881
15.9504 | -34.865 | -25.4961
15.9565 | -34.8655 | -25.5039
15.9626 | -34.866 | -25.5118
15.9687 | -34.8665 | -25.5194
15.9749 | -34.867 | -25.527
15.981 | -34.8674 | -25.5349
15.9871 | -34.8679 | -25.5428
15.9932 | -34.8685 | -25.5505
15.9993 | -34.869 | -25.5585
16.0054 | -34.8695 | -25.5658
16.0115 | -34.87 | -25.5736
16.0176 | -34.8705 | -25.5815
16.0237 | -34.871 | -25.5895
16.0298 | -34.8715 | -25.5973
16.0359 | -34.872 | -25.605
16.042 | -34.8725 | -25.6125
16.0481 | -34.873 | -25.6204
16.0542 | -34.8735 | -25.6283
16.0603 | -34.8741 | -25.636
16.0664 | -34.8745 | -25.6437
16.0725 | -34.8751 | -25.6513
16.0786 | -34.8756 | -25.6592
16.0847 | -34.8761 | -25.6671
16.0908 | -34.8766 | -25.6749
16.0969 | -34.8771 | -25.6827
16.103 | -34.8776 | -25.6906
16.1091 | -34.8781 | -25.6983
16.1152 | -34.8786 | -25.7058
16.1214 | -34.8791 | -25.7137
16.1275 | -34.8796 | -25.7216
16.1336 | -34.8801 | -25.7293
16.1397 | -34.8806 | -25.7369
16.1458 | -34.8811 | -25.7445
16.1519 | -34.8816 | -25.7525
16.158 | -34.8821 | -25.7602
16.1641 | -34.8826 | -25.7682
16.1702 | -34.8831 | -25.776
16.1763 | -34.8836 | -25.7837
16.1824 | -34.8841 | -25.7917
16.1885 | -34.8846 | -25.7993
16.1946 | -34.8851 | -25.8071
16.2007 | -34.8856 | -25.8148
16.2068 | -34.8862 | -25.8225
16.2129 | -34.8867 | -25.8304
16.219 | -34.8872 | -25.8378
16.2251 | -34.8878 | -25.8453
16.2312 | -34.8883 | -25.853
16.2373 | -34.8888 | -25.8606
16.2434 | -34.8893 | -25.8681
16.2495 | -34.8898 | -25.8757
16.2556 | -34.8904 | -25.8835
16.2618 | -34.8909 | -25.8912
16.2679 | -34.8914 | -25.8992
16.274 | -34.8919 | -25.907
16.2801 | -34.8924 | -25.9147
16.2862 | -34.8929 | -25.9222
16.2923 | -34.8934 | -25.93
16.2984 | -34.8939 | -25.9377
16.3045 | -34.8944 | -25.9455
16.3106 | -34.8949 | -25.9532
16.3167 | -34.8954 | -25.961
16.3228 | -34.8959 | -25.9685
16.3289 | -34.8964 | -25.9763
16.335 | -34.8969 | -25.9839
16.3411 | -34.8974 | -25.9917
16.3472 | -34.898 | -25.9995
16.3533 | -34.8985 | -26.0072
16.3594 | -34.899 | -26.0146
16.3655 | -34.8994 | -26.0225
16.3716 | -34.8999 | -26.03
16.3777 | -34.9005 | -26.0376
16.3838 | -34.901 | -26.0453
16.3899 | -34.9015 | -26.0527
16.396 | -34.902 | -26.0603
16.4021 | -34.9025 | -26.0681
16.4083 | -34.903 | -26.0758
16.4144 | -34.9035 | -26.0836
16.4205 | -34.9041 | -26.0913
16.4266 | -34.9046 | -26.0986
16.4327 | -34.905 | -26.1064
16.4388 | -34.9055 | -26.1142
16.4449 | -34.906 | -26.1218
16.451 | -34.9065 | -26.1296
16.4571 | -34.907 | -26.1374
16.4632 | -34.9075 | -26.145
16.4693 | -34.908 | -26.1528
16.4754 | -34.9085 | -26.1605
16.4815 | -34.909 | -26.1682
16.4876 | -34.9095 | -26.1758
16.4937 | -34.91 | -26.1836
16.4998 | -34.9105 | -26.1912
16.5059 | -34.9111 | -26.1989
16.512 | -34.9115 | -26.2067
16.5181 | -34.9121 | -26.2144
16.5242 | -34.9126 | -26.2222
16.5303 | -34.9131 | -26.2296
16.5364 | -34.9136 | -26.2372
16.5425 | -34.9141 | -26.245
16.5487 | -34.9146 | -26.2524
16.5548 | -34.9151 | -26.2599
16.5609 | -34.9156 | -26.2677
16.567 | -34.9161 | -26.2751
16.5731 | -34.9166 | -26.2829
16.5792 | -34.9171 | -26.2906
16.5853 | -34.9176 | -26.2981
16.5914 | -34.9181 | -26.306
16.5975 | -34.9186 | -26.3138
16.6036 | -34.9191 | -26.3212
16.6097 | -34.9196 | -26.329
16.6158 | -34.9201 | -26.3368
16.6219 | -34.9206 | -26.3444
16.628 | -34.9212 | -26.352
16.6341 | -34.9217 | -26.3598
16.6402 | -34.9222 | -26.3673
16.6463 | -34.9227 | -26.3749
16.6524 | -34.9232 | -26.3824
16.6585 | -34.9237 | -26.3901
16.6646 | -34.9242 | -26.3977
16.6707 | -34.9247 | -26.4051
16.6768 | -34.9252 | -26.4127
16.6829 | -34.9256 | -26.4205
16.689 | -34.9261 | -26.4281
16.6952 | -34.9266 | -26.4357
16.7013 | -34.9271 | -26.4432
16.7074 | -34.9276 | -26.4506
16.7135 | -34.9281 | -26.4582
16.7196 | -34.9286 | -26.4658
16.7257 | -34.9291 | -26.4735
16.7318 | -34.9296 | -26.481
16.7379 | -34.9301 | -26.4886
16.744 | -34.9306 | -26.4961
16.7501 | -34.9311 | -26.5039
16.7562 | -34.9316 | -26.5115
16.7623 | -34.9321 | -26.519
16.7684 | -34.9326 | -26.5266
16.7745 | -34.9331 | -26.5344
16.7806 | -34.9336 | -26.5415
16.7867 | -34.9341 | -26.549
16.7928 | -34.9347 | -26.5564
16.7989 | -34.9352 | -26.5638
16.805 | -34.9357 | -26.5714
16.8111 | -34.9362 | -26.5788
16.8172 | -34.9367 | -26.5864
16.8233 | -34.9372 | -26.594
16.8294 | -34.9377 | -26.6014
16.8356 | -34.9382 | -26.609
16.8417 | -34.9387 | -26.6165
16.8478 | -34.9391 | -26.624
16.8539 | -34.9395 | -26.6316
16.86 | -34.94 | -26.6393
16.8661 | -34.9405 | -26.6468
16.8722 | -34.9411 | -26.6546
16.8783 | -34.9415 | -26.6622
16.8844 | -34.942 | -26.6697
16.8905 | -34.9425 | -26.6774
16.8966 | -34.943 | -26.6851
16.9027 | -34.9435 | -26.6927
16.9088 | -34.944 | -26.7005
16.9149 | -34.9445 | -26.7082
16.921 | -34.945 | -26.7157
16.9271 | -34.9455 | -26.7234
16.9332 | -34.946 | -26.731
16.9393 | -34.9464 | -26.7387
16.9454 | -34.9469 | -26.7462
16.9515 | -34.9474 | -26.7541
16.9576 | -34.948 | -26.7614
16.9637 | -34.9485 | -26.7689
16.9698 | -34.949 | -26.7764
16.9759 | -34.9495 | -26.7838
16.9821 | -34.9501 | -26.7914
16.9882 | -34.9506 | -26.7986
16.9943 | -34.951 | -26.8062
17.0004 | -34.9515 | -26.8136
17.0065 | -34.9521 | -26.8212
17.0126 | -34.9526 | -26.8289
17.0187 | -34.953 | -26.8366
17.0248 | -34.9536 | -26.8441
17.0309 | -34.9541 | -26.8516
17.037 | -34.9546 | -26.859
17.0431 | -34.9551 | -26.8667
17.0492 | -34.9556 | -26.8743
17.0553 | -34.956 | -26.8819
17.0614 | -34.9565 | -26.8894
17.0675 | -34.957 | -26.897
17.0736 | -34.9575 | -26.9046
17.0797 | -34.958 | -26.912
17.0858 | -34.9585 | -26.9196
17.0919 | -34.959 | -26.9271
17.098 | -34.9595 | -26.9346
17.1041 | -34.96 | -26.942
17.1102 | -34.9605 | -26.9495
17.1163 | -34.961 | -26.9569
17.1225 | -34.9615 | -26.9644
17.1286 | -34.962 | -26.9719
17.1347 | -34.9625 | -26.9795
17.1408 | -34.963 | -26.9869
17.1469 | -34.9635 | -26.9944
17.153 | -34.9639 | -27.0019
17.1591 | -34.9644 | -27.0095
17.1652 | -34.9649 | -27.0167
17.1713 | -34.9654 | -27.0242
17.1774 | -34.9659 | -27.0316
17.1835 | -34.9664 | -27.0389
17.1896 | -34.9669 | -27.0464
17.1957 | -34.9674 | -27.0539
17.2018 | -34.9679 | -27.061
17.2079 | -34.9684 | -27.0686
17.214 | -34.9689 | -27.0762
17.2201 | -34.9693 | -27.0837
17.2262 | -34.9698 | -27.0913
17.2323 | -34.9704 | -27.0986
17.2384 | -34.9709 | -27.1059
17.2445 | -34.9714 | -27.1133
17.2506 | -34.9719 | -27.1209
17.2567 | -34.9723 | -27.1283
17.2628 | -34.9728 | -27.1358
17.269 | -34.9733 | -27.1432
17.2751 | -34.9738 | -27.1508
17.2812 | -34.9743 | -27.1584
17.2873 | -34.9748 | -27.1659
17.2934 | -34.9753 | -27.1735
17.2995 | -34.9758 | -27.181
17.3056 | -34.9763 | -27.1884
17.3117 | -34.9768 | -27.196
17.3178 | -34.9773 | -27.2036
17.3239 | -34.9778 | -27.2108
17.33 | -34.9782 | -27.2184
17.3361 | -34.9787 | -27.2258
17.3422 | -34.9792 | -27.233
17.3483 | -34.9797 | -27.2406
17.3544 | -34.9802 | -27.2481
17.3605 | -34.9807 | -27.2556
17.3666 | -34.9812 | -27.2629
17.3727 | -34.9817 | -27.2703
17.3788 | -34.9822 | -27.2777
17.3849 | -34.9827 | -27.2851
17.391 | -34.9832 | -27.2926
17.3971 | -34.9837 | -27.3001
17.4032 | -34.9842 | -27.3076
17.4094 | -34.9847 | -27.3149
17.4155 | -34.9852 | -27.3221
17.4216 | -34.9856 | -27.3295
17.4277 | -34.9861 | -27.3371
17.4338 | -34.9866 | -27.3445
17.4399 | -34.9871 | -27.3521
17.446 | -34.9876 | -27.3593
17.4521 | -34.9882 | -27.3668
17.4582 | -34.9887 | -27.3743
17.4643 | -34.9891 | -27.3816
17.4704 | -34.9896 | -27.389
17.4765 | -34.9901 | -27.3966
17.4826 | -34.9906 | -27.4039
17.4887 | -34.9911 | -27.4114
17.4948 | -34.9915 | -27.4189
17.5009 | -34.992 | -27.4263
17.507 | -34.9925 | -27.4338
17.5131 | -34.993 | -27.4412
17.5192 | -34.9935 | -27.4486
17.5253 | -34.994 | -27.4561
17.5314 | -34.9945 | -27.4634
17.5375 | -34.995 | -27.4708
17.5436 | -34.9955 | -27.4783
17.5497 | -34.996 | -27.4857
17.5559 | -34.9965 | -27.4932
17.562 | -34.997 | -27.5005
17.5681 | -34.9975 | -27.5078
17.5742 | -34.998 | -27.5153
17.5803 | -34.9985 | -27.5227
17.5864 | -34.999 | -27.5301
17.5925 | -34.9995 | -27.5376
17.5986 | -35 | -27.5451
17.6047 | -35.0005 | -27.5526
17.6108 | -35.001 | -27.56
17.6169 | -35.0015 | -27.5674
17.623 | -35.002 | -27.5748
17.6291 | -35.0024 | -27.5822
17.6352 | -35.0029 | -27.5896
17.6413 | -35.0034 | -27.5971
17.6474 | -35.0039 | -27.6045
17.6535 | -35.0044 | -27.6118
17.6596 | -35.0049 | -27.6191
17.6657 | -35.0054 | -27.6265
17.6718 | -35.0059 | -27.6338
17.6779 | -35.0064 | -27.6411
17.684 | -35.0069 | -27.6485
17.6901 | -35.0073 | -27.656
17.6963 | -35.0078 | -27.6633
17.7024 | -35.0083 | -27.6707
17.7085 | -35.0088 | -27.678
17.7146 | -35.0093 | -27.6853
17.7207 | -35.0098 | -27.6926
17.7268 | -35.0103 | -27.6998
17.7329 | -35.0108 | -27.7073
17.739 | -35.0113 | -27.7147
17.7451 | -35.0118 | -27.722
17.7512 | -35.0123 | -27.7295
17.7573 | -35.0128 | -27.7372
17.7634 | -35.0133 | -27.7443
17.7695 | -35.0138 | -27.7518
17.7756 | -35.0142 | -27.7593
17.7817 | -35.0147 | -27.7667
17.7878 | -35.0152 | -27.774
17.7939 | -35.0157 | -27.7814
17.8 | -35.0162 | -27.7886
17.8061 | -35.0167 | -27.7961
17.8122 | -35.0172 | -27.8034
17.8183 | -35.0177 | -27.8108
17.8244 | -35.0182 | -27.8181
17.8305 | -35.0187 | -27.8255
17.8366 | -35.0192 | -27.8326
17.8428 | -35.0197 | -27.8402
17.8489 | -35.0202 | -27.8475
17.855 | -35.0207 | -27.8548
17.8611 | -35.0211 | -27.8622
17.8672 | -35.0217 | -27.8695
17.8733 | -35.0221 | -27.8767
17.8794 | -35.0226 | -27.8841
17.8855 | -35.0231 | -27.8913
17.8916 | -35.0236 | -27.8988
17.8977 | -35.0241 | -27.9063
17.9038 | -35.0246 | -27.9134
17.9099 | -35.025 | -27.9208
17.916 | -35.0255 | -27.9283
17.9221 | -35.026 | -27.9358
17.9282 | -35.0265 | -27.943
17.9343 | -35.027 | -27.9502
17.9404 | -35.0274 | -27.9574
17.9465 | -35.0279 | -27.9648
17.9526 | -35.0284 | -27.9723
17.9587 | -35.0289 | -27.9797
17.9648 | -35.0293 | -27.9868
17.9709 | -35.0298 | -27.9942
17.977 | -35.0303 | -28.0016
17.9832 | -35.0307 | -28.0089
17.9893 | -35.0313 | -28.0165
17.9954 | -35.0317 | -28.024
18.0015 | -35.0322 | -28.0313
18.0076 | -35.0327 | -28.0385
18.0137 | -35.0331 | -28.0458
18.0198 | -35.0336 | -28.0531
18.0259 | -35.0341 | -28.0605
18.032 | -35.0345 | -28.0679
18.0381 | -35.035 | -28.0753
18.0442 | -35.0355 | -28.0826
18.0503 | -35.0359 | -28.0899
18.0564 | -35.0364 | -28.0971
18.0625 | -35.0369 | -28.1045
18.0686 | -35.0374 | -28.112
18.0747 | -35.0379 | -28.1194
18.0808 | -35.0384 | -28.1267
18.0869 | -35.0389 | -28.1341
18.093 | -35.0394 | -28.1415
18.0991 | -35.0398 | -28.1488
18.1052 | -35.0403 | -28.1562
18.1113 | -35.0408 | -28.1635
18.1174 | -35.0412 | -28.1707
18.1236 | -35.0418 | -28.1781
18.1297 | -35.0423 | -28.1855
18.1358 | -35.0427 | -28.193
18.1419 | -35.0433 | -28.2002
18.148 | -35.0438 | -28.2075
18.1541 | -35.0443 | -28.2147
18.1602 | -35.0448 | -28.2223
18.1663 | -35.0452 | -28.2295
18.1724 | -35.0457 | -28.2369
18.1785 | -35.0462 | -28.2441
18.1846 | -35.0467 | -28.2513
18.1907 | -35.0471 | -28.2587
18.1968 | -35.0476 | -28.2662
18.2029 | -35.0481 | -28.2735
18.209 | -35.0486 | -28.2808
18.2151 | -35.0491 | -28.2882
18.2212 | -35.0496 | -28.2952
18.2273 | -35.0501 | -28.3026
18.2334 | -35.0506 | -28.3099
18.2395 | -35.051 | -28.3173
18.2456 | -35.0515 | -28.3245
18.2517 | -35.0519 | -28.3318
18.2578 | -35.0524 | -28.3393
18.2639 | -35.0529 | -28.3466
18.2701 | -35.0534 | -28.3539
18.2762 | -35.0539 | -28.3611
18.2823 | -35.0544 | -28.3684
18.2884 | -35.0549 | -28.3756
18.2945 | -35.0554 | -28.3828
18.3006 | -35.0559 | -28.3901
18.3067 | -35.0564 | -28.3976
18.3128 | -35.0568 | -28.4051
18.3189 | -35.0573 | -28.4124
18.325 | -35.0577 | -28.4196
18.3311 | -35.0582 | -28.4267
18.3372 | -35.0587 | -28.4341
18.3433 | -35.0592 | -28.4414
18.3494 | -35.0597 | -28.4487
18.3555 | -35.0601 | -28.4559
18.3616 | -35.0606 | -28.463
18.3677 | -35.0611 | -28.4704
18.3738 | -35.0616 | -28.4777
18.3799 | -35.062 | -28.4851
18.386 | -35.0625 | -28.4925
18.3921 | -35.063 | -28.4999
18.3982 | -35.0634 | -28.5071
18.4043 | -35.0639 | -28.5143
18.4105 | -35.0644 | -28.5216
18.4166 | -35.0648 | -28.5288
18.4227 | -35.0653 | -28.5362
18.4288 | -35.0657 | -28.5434
18.4349 | -35.0662 | -28.5506
18.441 | -35.0666 | -28.5579
18.4471 | -35.0671 | -28.5652
18.4532 | -35.0676 | -28.5727
18.4593 | -35.0681 | -28.5802
18.4654 | -35.0685 | -28.5874
18.4715 | -35.069 | -28.5947
18.4776 | -35.0695 | -28.6023
18.4837 | -35.07 | -28.6097
18.4898 | -35.0704 | -28.6171
18.4959 | -35.0709 | -28.6245
18.502 | -35.0714 | -28.6316
18.5081 | -35.0719 | -28.6392
18.5142 | -35.0723 | -28.6465
18.5203 | -35.0728 | -28.6538
18.5264 | -35.0733 | -28.661
18.5325 | -35.0737 | -28.6682
18.5386 | -35.0743 | -28.6753
18.5447 | -35.0747 | -28.6826
18.5508 | -35.0752 | -28.69
18.557 | -35.0757 | -28.6974
18.5631 | -35.0762 | -28.7047
18.5692 | -35.0767 | -28.7117
18.5753 | -35.0771 | -28.7188
18.5814 | -35.0776 | -28.7261
18.5875 | -35.0781 | -28.7334
18.5936 | -35.0786 | -28.7406
18.5997 | -35.0791 | -28.748
18.6058 | -35.0795 | -28.7554
18.6119 | -35.08 | -28.7629
18.618 | -35.0805 | -28.7701
18.6241 | -35.081 | -28.7775
18.6302 | -35.0815 | -28.7849
18.6363 | -35.0819 | -28.7924
18.6424 | -35.0824 | -28.7997
18.6485 | -35.0829 | -28.807
18.6546 | -35.0833 | -28.8142
18.6607 | -35.0838 | -28.8214
18.6668 | -35.0843 | -28.8287
18.6729 | -35.0848 | -28.8362
18.679 | -35.0853 | -28.8432
18.6851 | -35.0858 | -28.8504
18.6912 | -35.0862 | -28.8578
18.6974 | -35.0867 | -28.865
18.7035 | -35.0872 | -28.8724
18.7096 | -35.0876 | -28.8799
18.7157 | -35.0881 | -28.887
18.7218 | -35.0886 | -28.8943
18.7279 | -35.0891 | -28.9017
18.734 | -35.0896 | -28.909
18.7401 | -35.09 | -28.9162
18.7462 | -35.0905 | -28.9229
18.7523 | -35.091 | -28.9302
18.7584 | -35.0914 | -28.9376
18.7645 | -35.0919 | -28.9449
18.7706 | -35.0923 | -28.9521
18.7767 | -35.0928 | -28.9594
18.7828 | -35.0933 | -28.9668
18.7889 | -35.0937 | -28.9743
18.795 | -35.0942 | -28.9814
18.8011 | -35.0947 | -28.9888
18.8072 | -35.0951 | -28.9962
18.8133 | -35.0956 | -29.0033
18.8194 | -35.096 | -29.0103
18.8255 | -35.0965 | -29.0175
18.8316 | -35.097 | -29.0249
18.8377 | -35.0975 | -29.0319
18.8439 | -35.098 | -29.0393
18.85 | -35.0984 | -29.0466
18.8561 | -35.0989 | -29.0538
18.8622 | -35.0994 | -29.0611
18.8683 | -35.0999 | -29.0685
18.8744 | -35.1004 | -29.0756
18.8805 | -35.1009 | -29.0828
18.8866 | -35.1013 | -29.09
18.8927 | -35.1018 | -29.0971
18.8988 | -35.1022 | -29.1042
18.9049 | -35.1027 | -29.1113
18.911 | -35.1032 | -29.1185
18.9171 | -35.1037 | -29.1255
18.9232 | -35.1041 | -29.1327
18.9293 | -35.1046 | -29.1398
18.9354 | -35.105 | -29.147
18.9415 | -35.1055 | -29.1542
18.9476 | -35.106 | -29.1613
18.9537 | -35.1065 | -29.1688
18.9598 | -35.1069 | -29.176
18.9659 | -35.1074 | -29.183
18.972 | -35.1079 | -29.1902
18.9781 | -35.1084 | -29.1975
18.9843 | -35.1088 | -29.2048
18.9904 | -35.1093 | -29.2121
18.9965 | -35.1098 | -29.2192
19.0026 | -35.1103 | -29.2265
19.0087 | -35.1107 | -29.2338
19.0148 | -35.1112 | -29.241
19.0209 | -35.1117 | -29.2483
19.027 | -35.1122 | -29.2556
19.0331 | -35.1126 | -29.2626
19.0392 | -35.1131 | -29.2701
19.0453 | -35.1136 | -29.2773
19.0514 | -35.1141 | -29.2847
19.0575 | -35.1146 | -29.292
19.0636 | -35.115 | -29.2988
19.0697 | -35.1155 | -29.3059
19.0758 | -35.116 | -29.313
19.0819 | -35.1165 | -29.3203
19.088 | -35.117 | -29.3277
19.0941 | -35.1175 | -29.335
19.1002 | -35.1179 | -29.342
19.1063 | -35.1184 | -29.3491
19.1124 | -35.1189 | -29.3563
19.1185 | -35.1193 | -29.3634
19.1246 | -35.1198 | -29.3707
19.1308 | -35.1203 | -29.3778
19.1369 | -35.1208 | -29.3848
19.143 | -35.1213 | -29.3919
19.1491 | -35.1217 | -29.3992
19.1552 | -35.1222 | -29.4067
19.1613 | -35.1227 | -29.4139
19.1674 | -35.1231 | -29.4212
19.1735 | -35.1236 | -29.4283
19.1796 | -35.124 | -29.4355
19.1857 | -35.1245 | -29.4427
19.1918 | -35.125 | -29.4498
19.1979 | -35.1255 | -29.4567
19.204 | -35.126 | -29.4638
19.2101 | -35.1265 | -29.4709
19.2162 | -35.127 | -29.4782
19.2223 | -35.1275 | -29.4853
19.2284 | -35.1279 | -29.4927
19.2345 | -35.1284 | -29.5
19.2406 | -35.1288 | -29.5069
19.2467 | -35.1293 | -29.5142
19.2528 | -35.1298 | -29.5213
19.2589 | -35.1302 | -29.5284
19.265 | -35.1307 | -29.5359
19.2712 | -35.1311 | -29.543
19.2773 | -35.1315 | -29.5499
19.2834 | -35.132 | -29.557
19.2895 | -35.1324 | -29.5642
19.2956 | -35.1329 | -29.5715
19.3017 | -35.1334 | -29.5788
19.3078 | -35.1339 | -29.5862
19.3139 | -35.1344 | -29.593
19.32 | -35.1348 | -29.5999
19.3261 | -35.1352 | -29.607
19.3322 | -35.1357 | -29.6139
19.3383 | -35.1362 | -29.621
19.3444 | -35.1366 | -29.6277
19.3505 | -35.1371 | -29.6351
19.3566 | -35.1375 | -29.6422
19.3627 | -35.138 | -29.6493
19.3688 | -35.1385 | -29.6563
19.3749 | -35.1389 | -29.6636
19.381 | -35.1394 | -29.6704
19.3871 | -35.1399 | -29.6774
19.3932 | -35.1403 | -29.6846
19.3993 | -35.1407 | -29.6917
19.4054 | -35.1412 | -29.6989
19.4115 | -35.1417 | -29.706
19.4177 | -35.1421 | -29.7132
19.4238 | -35.1426 | -29.7204
19.4299 | -35.1431 | -29.7274
19.436 | -35.1436 | -29.7347
19.4421 | -35.1441 | -29.7414
19.4482 | -35.1446 | -29.7484
19.4543 | -35.145 | -29.7548
19.4604 | -35.1454 | -29.7618
19.4665 | -35.1459 | -29.7687
19.4726 | -35.1464 | -29.7759
19.4787 | -35.1468 | -29.7829
19.4848 | -35.1472 | -29.79
19.4909 | -35.1477 | -29.7972
19.497 | -35.1481 | -29.8041
19.5031 | -35.1486 | -29.8109
19.5092 | -35.1491 | -29.8182
19.5153 | -35.1496 | -29.8253
19.5214 | -35.15 | -29.8325
19.5275 | -35.1504 | -29.8393
19.5336 | -35.1509 | -29.8461
19.5397 | -35.1513 | -29.8531
19.5458 | -35.1518 | -29.86
19.5519 | -35.1523 | -29.8668
19.5581 | -35.1527 | -29.8737
19.5642 | -35.1532 | -29.8808
19.5703 | -35.1536 | -29.8877
19.5764 | -35.1541 | -29.8947
19.5825 | -35.1546 | -29.9016
19.5886 | -35.1551 | -29.9088
19.5947 | -35.1556 | -29.9156
19.6008 | -35.156 | -29.9226
19.6069 | -35.1564 | -29.9299
19.613 | -35.1569 | -29.9372
19.6191 | -35.1574 | -29.9445
19.6252 | -35.1578 | -29.9516
19.6313 | -35.1583 | -29.9578
19.6374 | -35.1587 | -29.9646
19.6435 | -35.1592 | -29.9716
19.6496 | -35.1597 | -29.9787
19.6557 | -35.1601 | -29.9856
19.6618 | -35.1606 | -29.992
19.6679 | -35.1611 | -29.999
19.674 | -35.1616 | -30.0061
19.6801 | -35.162 | -30.0127
19.6862 | -35.1624 | -30.0199
19.6923 | -35.1629 | -30.0266
19.6984 | -35.1633 | -30.0334
19.7046 | -35.1638 | -30.04
19.7107 | -35.1643 | -30.0467
19.7168 | -35.1648 | -30.0535
19.7229 | -35.1652 | -30.0606
19.729 | -35.1657 | -30.0676
19.7351 | -35.1661 | -30.0745
19.7412 | -35.1666 | -30.0815
19.7473 | -35.167 | -30.0885
19.7534 | -35.1675 | -30.0958
19.7595 | -35.168 | -30.1027
19.7656 | -35.1684 | -30.1098
19.7717 | -35.1689 | -30.1169
19.7778 | -35.1693 | -30.1237
19.7839 | -35.1698 | -30.1306
19.79 | -35.1703 | -30.1376
19.7961 | -35.1707 | -30.1446
19.8022 | -35.1712 | -30.151
19.8083 | -35.1716 | -30.1581
19.8144 | -35.1721 | -30.1648
19.8205 | -35.1726 | -30.1719
19.8266 | -35.1731 | -30.1785
19.8327 | -35.1735 | -30.1852
19.8388 | -35.174 | -30.192
19.845 | -35.1745 | -30.1987
19.8511 | -35.1749 | -30.2057
19.8572 | -35.1753 | -30.2127
19.8633 | -35.1758 | -30.2197
19.8694 | -35.1763 | -30.2267
19.8755 | -35.1767 | -30.2331
19.8816 | -35.1772 | -30.2402
19.8877 | -35.1776 | -30.2471
19.8938 | -35.1781 | -30.2541
19.8999 | -35.1785 | -30.261
19.906 | -35.179 | -30.2682
19.9121 | -35.1795 | -30.2749
19.9182 | -35.18 | -30.2817
19.9243 | -35.1804 | -30.2884
19.9304 | -35.1808 | -30.2952
19.9365 | -35.1812 | -30.3021
19.9426 | -35.1817 | -30.3087
19.9487 | -35.1822 | -30.3157
19.9548 | -35.1826 | -30.323
19.9609 | -35.1831 | -30.3299
19.967 | -35.1836 | -30.3366
19.9731 | -35.184 | -30.3432
19.9792 | -35.1845 | -30.3498
19.9853 | -35.185 | -30.3565
19.9915 | -35.1854 | -30.3634
19.9976 | -35.1859 | -30.3702
20.0037 | -35.1863 | -30.3768
20.0098 | -35.1868 | -30.3833
20.0159 | -35.1873 | -30.3901
20.022 | -35.1878 | -30.3968
20.0281 | -35.1882 | -30.4037
20.0342 | -35.1887 | -30.4109
20.0403 | -35.1891 | -30.4178
20.0464 | -35.1896 | -30.4252
20.0525 | -35.1901 | -30.4317
20.0586 | -35.1906 | -30.4388
20.0647 | -35.191 | -30.4459
20.0708 | -35.1915 | -30.4528
20.0769 | -35.192 | -30.4597
20.083 | -35.1924 | -30.4664
20.0891 | -35.1929 | -30.4732
20.0952 | -35.1933 | -30.4799
20.1013 | -35.1937 | -30.4868
20.1074 | -35.1941 | -30.4935
20.1135 | -35.1946 | -30.5006
20.1196 | -35.195 | -30.5068
20.1257 | -35.1955 | -30.514
20.1319 | -35.1959 | -30.5207
20.138 | -35.1964 | -30.5275
20.1441 | -35.1969 | -30.5346
20.1502 | -35.1973 | -30.5414
20.1563 | -35.1977 | -30.5477
20.1624 | -35.1982 | -30.5545
20.1685 | -35.1987 | -30.5613
20.1746 | -35.1991 | -30.5682
20.1807 | -35.1996 | -30.5749
20.1868 | -35.2 | -30.5819
20.1929 | -35.2005 | -30.5883
20.199 | -35.2009 | -30.5949
20.2051 | -35.2014 | -30.6018
20.2112 | -35.2018 | -30.6085
20.2173 | -35.2023 | -30.6153
20.2234 | -35.2027 | -30.6219
20.2295 | -35.2033 | -30.6283
20.2356 | -35.2037 | -30.6348
20.2417 | -35.2042 | -30.6416
20.2478 | -35.2046 | -30.6484
20.2539 | -35.2051 | -30.6554
20.26 | -35.2055 | -30.6619
20.2661 | -35.206 | -30.6685
20.2723 | -35.2065 | -30.6753
20.2784 | -35.2069 | -30.6822
20.2845 | -35.2074 | -30.6889
20.2906 | -35.2079 | -30.6959
20.2967 | -35.2084 | -30.7023
20.3028 | -35.2089 | -30.7093
20.3089 | -35.2094 | -30.716
20.315 | -35.2099 | -30.7229
20.3211 | -35.2103 | -30.7295
20.3272 | -35.2107 | -30.7365
20.3333 | -35.2112 | -30.7426
20.3394 | -35.2117 | -30.7491
20.3455 | -35.2121 | -30.7559
20.3516 | -35.2126 | -30.7624
20.3577 | -35.213 | -30.7688
20.3638 | -35.2135 | -30.7753
20.3699 | -35.2139 | -30.7817
20.376 | -35.2144 | -30.7883
20.3821 | -35.2149 | -30.795
20.3882 | -35.2153 | -30.8023
20.3943 | -35.2157 | -30.8086
20.4004 | -35.2162 | -30.8153
20.4065 | -35.2166 | -30.8219
20.4126 | -35.2171 | -30.8286
20.4188 | -35.2175 | -30.8357
20.4249 | -35.2179 | -30.8424
20.431 | -35.2183 | -30.849
20.4371 | -35.2188 | -30.8552
20.4432 | -35.2193 | -30.8617
20.4493 | -35.2197 | -30.8685
20.4554 | -35.2201 | -30.8752
20.4615 | -35.2206 | -30.8824
20.4676 | -35.2211 | -30.8892
20.4737 | -35.2216 | -30.8955
20.4798 | -35.2221 | -30.9021
20.4859 | -35.2225 | -30.909
20.492 | -35.223 | -30.9156
20.4981 | -35.2234 | -30.9224
20.5042 | -35.2239 | -30.9292
20.5103 | -35.2244 | -30.9354
20.5164 | -35.2248 | -30.9421
20.5225 | -35.2252 | -30.9486
20.5286 | -35.2257 | -30.9551
20.5347 | -35.2261 | -30.9618
20.5408 | -35.2266 | -30.9679
20.5469 | -35.2271 | -30.9746
20.553 | -35.2275 | -30.9809
20.5592 | -35.228 | -30.9874
20.5653 | -35.2285 | -30.9939
20.5714 | -35.2289 | -31.0001
20.5775 | -35.2294 | -31.0067
20.5836 | -35.2299 | -31.0135
20.5897 | -35.2303 | -31.0205
20.5958 | -35.2307 | -31.0273
20.6019 | -35.2312 | -31.0341
20.608 | -35.2317 | -31.0409
20.6141 | -35.2321 | -31.0475
20.6202 | -35.2326 | -31.0546
20.6263 | -35.2331 | -31.0612
20.6324 | -35.2336 | -31.0677
20.6385 | -35.234 | -31.0744
20.6446 | -35.2344 | -31.0811
20.6507 | -35.2349 | -31.0875
20.6568 | -35.2353 | -31.094
20.6629 | -35.2358 | -31.1007
20.669 | -35.2362 | -31.1075
20.6751 | -35.2366 | -31.1141
20.6812 | -35.2371 | -31.1202
20.6873 | -35.2375 | -31.127
20.6934 | -35.238 | -31.1336
20.6995 | -35.2384 | -31.14
20.7057 | -35.2389 | -31.1465
20.7118 | -35.2393 | -31.1531
20.7179 | -35.2399 | -31.1591
20.724 | -35.2403 | -31.1655
20.7301 | -35.2408 | -31.1716
20.7362 | -35.2412 | -31.1785
20.7423 | -35.2417 | -31.1846
20.7484 | -35.2422 | -31.1914
20.7545 | -35.2426 | -31.197
20.7606 | -35.243 | -31.2034
20.7667 | -35.2435 | -31.2097
20.7728 | -35.2439 | -31.216
20.7789 | -35.2444 | -31.2226
20.785 | -35.2449 | -31.2291
20.7911 | -35.2454 | -31.235
20.7972 | -35.2458 | -31.2417
20.8033 | -35.2462 | -31.2484
20.8094 | -35.2467 | -31.2549
20.8155 | -35.2472 | -31.2612
20.8216 | -35.2476 | -31.2678
20.8277 | -35.2481 | -31.2733
20.8338 | -35.2486 | -31.28
20.8399 | -35.2491 | -31.2865
20.8461 | -35.2496 | -31.2929
20.8522 | -35.2501 | -31.2998
20.8583 | -35.2506 | -31.306
20.8644 | -35.251 | -31.3126
20.8705 | -35.2515 | -31.3187
20.8766 | -35.2519 | -31.3254
20.8827 | -35.2524 | -31.3317
20.8888 | -35.2528 | -31.3383
20.8949 | -35.2532 | -31.3449
20.901 | -35.2536 | -31.351
20.9071 | -35.2541 | -31.3577
20.9132 | -35.2545 | -31.3643
20.9193 | -35.255 | -31.371
20.9254 | -35.2554 | -31.3773
20.9315 | -35.2558 | -31.3835
20.9376 | -35.2562 | -31.39
20.9437 | -35.2567 | -31.3962
20.9498 | -35.2571 | -31.4025
20.9559 | -35.2576 | -31.4097
20.962 | -35.2582 | -31.4157
20.9681 | -35.2587 | -31.4215
20.9742 | -35.2592 | -31.4286
20.9803 | -35.2597 | -31.4346
20.9864 | -35.2601 | -31.4411
20.9926 | -35.2606 | -31.4475
20.9987 | -35.2611 | -31.4539
21.0048 | -35.2614 | -31.4606
21.0109 | -35.262 | -31.4674
21.017 | -35.2624 | -31.4742
21.0231 | -35.2629 | -31.4805
21.0292 | -35.2635 | -31.487
21.0353 | -35.2639 | -31.4932
21.0414 | -35.2644 | -31.5
21.0475 | -35.2647 | -31.5066
21.0536 | -35.2652 | -31.5127
21.0597 | -35.2657 | -31.5191
21.0658 | -35.2661 | -31.5252
21.0719 | -35.2666 | -31.5314
21.078 | -35.2671 | -31.5381
21.0841 | -35.2675 | -31.5443
21.0902 | -35.268 | -31.5509
21.0963 | -35.2684 | -31.5575
21.1024 | -35.2688 | -31.564
21.1085 | -35.2692 | -31.5707
21.1146 | -35.2696 | -31.5765
21.1207 | -35.27 | -31.5832
21.1268 | -35.2705 | -31.5892
21.133 | -35.2711 | -31.5958
21.1391 | -35.2716 | -31.6011
21.1452 | -35.272 | -31.6076
21.1513 | -35.2725 | -31.6136
21.1574 | -35.273 | -31.6198
21.1635 | -35.2735 | -31.6262
21.1696 | -35.2739 | -31.6323
21.1757 | -35.2744 | -31.6376
21.1818 | -35.2748 | -31.6441
21.1879 | -35.2753 | -31.6509
21.194 | -35.2757 | -31.6573
21.2001 | -35.2761 | -31.6635
21.2062 | -35.2766 | -31.67
21.2123 | -35.2771 | -31.6764
21.2184 | -35.2775 | -31.6812
21.2245 | -35.278 | -31.6885
21.2306 | -35.2785 | -31.6943
21.2367 | -35.279 | -31.7003
21.2428 | -35.2793 | -31.7064
21.2489 | -35.2798 | -31.7117
21.255 | -35.2802 | -31.7178
21.2611 | -35.2806 | -31.7246
21.2672 | -35.2811 | -31.7312
21.2733 | -35.2815 | -31.7371
21.2795 | -35.2819 | -31.7427
21.2856 | -35.2823 | -31.7493
21.2917 | -35.2829 | -31.7555
21.2978 | -35.2834 | -31.7614
21.3039 | -35.284 | -31.7675
21.31 | -35.2844 | -31.7736
21.3161 | -35.2846 | -31.7796
21.3222 | -35.285 | -31.7852
21.3283 | -35.2854 | -31.791
21.3344 | -35.2859 | -31.7974
21.3405 | -35.2862 | -31.8038
21.3466 | -35.2868 | -31.8098
21.3527 | -35.2873 | -31.8152
21.3588 | -35.2878 | -31.8215
21.3649 | -35.2882 | -31.8277
21.371 | -35.2886 | -31.8342
21.3771 | -35.2892 | -31.8399
21.3832 | -35.2898 | -31.8463
21.3893 | -35.2901 | -31.8506
21.3954 | -35.2906 | -31.8569
21.4015 | -35.2911 | -31.8621
21.4076 | -35.2916 | -31.8686
21.4137 | -35.292 | -31.8753
21.4199 | -35.2924 | -31.8816
21.426 | -35.2928 | -31.8875
21.4321 | -35.2933 | -31.8931
21.4382 | -35.2938 | -31.8991
21.4443 | -35.2943 | -31.9053
21.4504 | -35.2947 | -31.9116
21.4565 | -35.2952 | -31.9167
21.4626 | -35.2957 | -31.923
21.4687 | -35.2962 | -31.9293
21.4748 | -35.2967 | -31.9348
21.4809 | -35.2972 | -31.941
21.487 | -35.2975 | -31.9472
21.4931 | -35.298 | -31.9519
21.4992 | -35.2984 | -31.9587
21.5053 | -35.2989 | -31.9647
21.5114 | -35.2994 | -31.9705
21.5175 | -35.2999 | -31.976
21.5236 | -35.3002 | -31.9821
21.5297 | -35.3006 | -31.9881
21.5358 | -35.301 | -31.9947
21.5419 | -35.3015 | -32.0014
21.548 | -35.3019 | -32.0078
21.5541 | -35.3024 | -32.0145
21.5602 | -35.3029 | -32.0209
21.5664 | -35.3034 | -32.0265
21.5725 | -35.3039 | -32.0331
21.5786 | -35.3043 | -32.0397
21.5847 | -35.3048 | -32.0456
21.5908 | -35.3051 | -32.0516
21.5969 | -35.3056 | -32.0582
21.603 | -35.3061 | -32.0646
21.6091 | -35.3066 | -32.071
21.6152 | -35.307 | -32.0764
21.6213 | -35.3076 | -32.0825
21.6274 | -35.3079 | -32.0889
21.6335 | -35.3084 | -32.0942
21.6396 | -35.3088 | -32.1003
21.6457 | -35.3092 | -32.106
21.6518 | -35.3096 | -32.1124
21.6579 | -35.3102 | -32.1188
21.664 | -35.3107 | -32.1258
21.6701 | -35.3112 | -32.1318
21.6762 | -35.3116 | -32.1383
21.6823 | -35.3121 | -32.1437
21.6884 | -35.3126 | -32.1501
21.6945 | -35.3132 | -32.1559
21.7006 | -35.3137 | -32.1625
21.7068 | -35.3141 | -32.1688
21.7129 | -35.3145 | -32.1749
21.719 | -35.315 | -32.1812
21.7251 | -35.3155 | -32.1874
21.7312 | -35.3159 | -32.1934
21.7373 | -35.3163 | -32.1982
21.7434 | -35.3167 | -32.2049
21.7495 | -35.3173 | -32.2108
21.7556 | -35.3177 | -32.217
21.7617 | -35.3181 | -32.2235
21.7678 | -35.3186 | -32.2298
21.7739 | -35.3191 | -32.2353
21.78 | -35.3196 | -32.2415
21.7861 | -35.3201 | -32.2476
21.7922 | -35.3205 | -32.2539
21.7983 | -35.3209 | -32.2597
21.8044 | -35.3214 | -32.2657
21.8105 | -35.3218 | -32.2709
21.8166 | -35.3223 | -32.2772
21.8227 | -35.3228 | -32.2834
21.8288 | -35.3232 | -32.2899
21.8349 | -35.3236 | -32.2962
21.841 | -35.3241 | -32.3017
21.8471 | -35.3247 | -32.3071
21.8533 | -35.3252 | -32.3127
21.8594 | -35.3256 | -32.3189
21.8655 | -35.326 | -32.3248
21.8716 | -35.3264 | -32.3304
21.8777 | -35.3269 | -32.337
21.8838 | -35.3273 | -32.3432
21.8899 | -35.3278 | -32.3487
21.896 | -35.3283 | -32.354
21.9021 | -35.3286 | -32.3602
21.9082 | -35.3292 | -32.3672
21.9143 | -35.3297 | -32.3731
21.9204 | -35.3301 | -32.3788
21.9265 | -35.3305 | -32.3845
21.9326 | -35.3309 | -32.3897
21.9387 | -35.3315 | -32.3956
21.9448 | -35.3321 | -32.4012
21.9509 | -35.3326 | -32.4068
21.957 | -35.333 | -32.4128
21.9631 | -35.3334 | -32.4185
21.9692 | -35.3339 | -32.4246
21.9753 | -35.3345 | -32.4304
21.9814 | -35.3349 | -32.436
21.9875 | -35.3353 | -32.442
21.9937 | -35.3357 | -32.4477
21.9998 | -35.3362 | -32.4538
22.0059 | -35.3367 | -32.4598
22.012 | -35.337 | -32.4658
22.0181 | -35.3375 | -32.4718
22.0242 | -35.338 | -32.478
22.0303 | -35.3384 | -32.4842
22.0364 | -35.3389 | -32.4904
22.0425 | -35.3393 | -32.4963
22.0486 | -35.3397 | -32.5025
22.0547 | -35.3401 | -32.5081
22.0608 | -35.3405 | -32.5137
22.0669 | -35.3409 | -32.5198
22.073 | -35.3414 | -32.5257
22.0791 | -35.3419 | -32.5317
22.0852 | -35.3422 | -32.5373
22.0913 | -35.3427 | -32.5433
22.0974 | -35.3431 | -32.5495
22.1035 | -35.3435 | -32.5557
22.1096 | -35.344 | -32.5617
22.1157 | -35.3444 | -32.5677
22.1218 | -35.3449 | -32.5733
22.1279 | -35.3453 | -32.5792
22.134 | -35.3458 | -32.5854
22.1402 | -35.3462 | -32.5911
22.1463 | -35.3467 | -32.5972
22.1524 | -35.3473 | -32.6032
22.1585 | -35.3477 | -32.6092
22.1646 | -35.3481 | -32.6152
22.1707 | -35.3485 | -32.6208
22.1768 | -35.3489 | -32.6267
22.1829 | -35.3494 | -32.6328
22.189 | -35.3498 | -32.6388
22.1951 | -35.3501 | -32.6444
22.2012 | -35.3505 | -32.6507
22.2073 | -35.3509 | -32.6568
22.2134 | -35.3514 | -32.6625
22.2195 | -35.3519 | -32.6685
22.2256 | -35.3523 | -32.6748
22.2317 | -35.3528 | -32.6805
22.2378 | -35.3533 | -32.6861
22.2439 | -35.3538 | -32.692
22.25 | -35.3542 | -32.6978
22.2561 | -35.3546 | -32.704
22.2622 | -35.355 | -32.7104
22.2683 | -35.3555 | -32.7161
22.2744 | -35.3561 | -32.7217
22.2806 | -35.3566 | -32.7278
22.2867 | -35.357 | -32.7339
22.2928 | -35.3576 | -32.7402
22.2989 | -35.358 | -32.7461
22.305 | -35.3584 | -32.7516
22.3111 | -35.3588 | -32.7568
22.3172 | -35.3591 | -32.7631
22.3233 | -35.3596 | -32.7691
22.3294 | -35.3601 | -32.775
22.3355 | -35.3604 | -32.7805
22.3416 | -35.3609 | -32.7865
22.3477 | -35.3613 | -32.7926
22.3538 | -35.3617 | -32.7983
22.3599 | -35.3622 | -32.8042
22.366 | -35.3626 | -32.8102
22.3721 | -35.3632 | -32.8158
22.3782 | -35.3637 | -32.8215
22.3843 | -35.3642 | -32.8274
22.3904 | -35.3646 | -32.833
22.3965 | -35.3651 | -32.839
22.4026 | -35.3656 | -32.8452
22.4087 | -35.366 | -32.8506
22.4148 | -35.3664 | -32.8565
22.4209 | -35.3669 | -32.8628
22.4271 | -35.3673 | -32.8687
22.4332 | -35.3677 | -32.8741
22.4393 | -35.3681 | -32.8802
22.4454 | -35.3686 | -32.8861
22.4515 | -35.3691 | -32.8918
22.4576 | -35.3695 | -32.8977
22.4637 | -35.3699 | -32.9037
22.4698 | -35.3704 | -32.9099
22.4759 | -35.3709 | -32.9155
22.482 | -35.3713 | -32.9213
22.4881 | -35.3717 | -32.9274
22.4942 | -35.3722 | -32.9331
22.5003 | -35.3726 | -32.9387
22.5064 | -35.3731 | -32.9444
22.5125 | -35.3734 | -32.9497
22.5186 | -35.3739 | -32.9555
22.5247 | -35.3744 | -32.9614
22.5308 | -35.3748 | -32.9674
22.5369 | -35.3753 | -32.9736
22.543 | -35.3757 | -32.9798
22.5491 | -35.3761 | -32.9862
22.5552 | -35.3765 | -32.9913
22.5613 | -35.377 | -32.9977
22.5675 | -35.3774 | -33.0036
22.5736 | -35.3779 | -33.0092
22.5797 | -35.3784 | -33.015
22.5858 | -35.3788 | -33.0209
22.5919 | -35.3793 | -33.027
22.598 | -35.3798 | -33.0329
22.6041 | -35.3802 | -33.0389
22.6102 | -35.3806 | -33.0449
22.6163 | -35.381 | -33.0503
22.6224 | -35.3815 | -33.0555
22.6285 | -35.3819 | -33.0607
22.6346 | -35.3824 | -33.0664
22.6407 | -35.3828 | -33.0719
22.6468 | -35.3832 | -33.0778
22.6529 | -35.3837 | -33.0828
22.659 | -35.3841 | -33.0884
22.6651 | -35.3845 | -33.0945
22.6712 | -35.385 | -33.1004
22.6773 | -35.3854 | -33.1057
22.6834 | -35.3859 | -33.1113
22.6895 | -35.3863 | -33.1172
22.6956 | -35.3867 | -33.1232
22.7017 | -35.3872 | -33.1289
22.7079 | -35.3877 | -33.1347
22.714 | -35.3882 | -33.1408
22.7201 | -35.3886 | -33.1465
22.7262 | -35.389 | -33.1524
22.7323 | -35.3894 | -33.1578
22.7384 | -35.3899 | -33.1637
22.7445 | -35.3903 | -33.1696
22.7506 | -35.3907 | -33.1751
22.7567 | -35.3912 | -33.1809
22.7628 | -35.3917 | -33.1864
22.7689 | -35.3921 | -33.1922
22.775 | -35.3926 | -33.1981
22.7811 | -35.393 | -33.2038
22.7872 | -35.3935 | -33.2092
22.7933 | -35.3939 | -33.2151
22.7994 | -35.3944 | -33.2211
22.8055 | -35.3948 | -33.227
22.8116 | -35.3953 | -33.2327
22.8177 | -35.3957 | -33.2387
22.8238 | -35.3962 | -33.2445
22.8299 | -35.3966 | -33.2498
22.836 | -35.3971 | -33.2552
22.8421 | -35.3975 | -33.2606
22.8482 | -35.3979 | -33.2664
22.8544 | -35.3984 | -33.2721
22.8605 | -35.3988 | -33.2779
22.8666 | -35.3993 | -33.2837
22.8727 | -35.3997 | -33.2895
22.8788 | -35.4002 | -33.2949
22.8849 | -35.4006 | -33.3004
22.891 | -35.401 | -33.3064
22.8971 | -35.4015 | -33.3123
22.9032 | -35.4019 | -33.3185
22.9093 | -35.4024 | -33.3245
22.9154 | -35.4028 | -33.3303
22.9215 | -35.4032 | -33.3359
22.9276 | -35.4037 | -33.342
22.9337 | -35.4042 | -33.3478
22.9398 | -35.4047 | -33.353
22.9459 | -35.4051 | -33.3587
22.952 | -35.4056 | -33.3644
22.9581 | -35.4061 | -33.3699
22.9642 | -35.4065 | -33.3755
22.9703 | -35.4069 | -33.3812
22.9764 | -35.4074 | -33.3871
22.9825 | -35.4078 | -33.3927
22.9886 | -35.4083 | -33.3987
22.9948 | -35.4087 | -33.4046
23.0009 | -35.4092 | -33.4104
23.007 | -35.4096 | -33.4162
23.0131 | -35.4101 | -33.4217
23.0192 | -35.4106 | -33.4274
23.0253 | -35.411 | -33.4333
23.0314 | -35.4114 | -33.4392
23.0375 | -35.4118 | -33.4452
23.0436 | -35.4124 | -33.4507
23.0497 | -35.4128 | -33.4562
23.0558 | -35.4133 | -33.4622
23.0619 | -35.4137 | -33.4678
23.068 | -35.4142 | -33.4738
23.0741 | -35.4146 | -33.4793
23.0802 | -35.4151 | -33.4852
23.0863 | -35.4155 | -33.4912
23.0924 | -35.416 | -33.497
23.0985 | -35.4164 | -33.5033
23.1046 | -35.4168 | -33.509
23.1107 | -35.4172 | -33.5144
23.1168 | -35.4177 | -33.5199
23.1229 | -35.4181 | -33.5256
23.129 | -35.4185 | -33.5312
23.1351 | -35.419 | -33.5371
23.1413 | -35.4193 | -33.5425
23.1474 | -35.4198 | -33.5484
23.1535 | -35.4202 | -33.5543
23.1596 | -35.4206 | -33.5602
23.1657 | -35.4211 | -33.5656
23.1718 | -35.4216 | -33.5709
23.1779 | -35.422 | -33.5766
23.184 | -35.4224 | -33.5819
23.1901 | -35.4228 | -33.5879
23.1962 | -35.4233 | -33.5935
23.2023 | -35.4237 | -33.5996
23.2084 | -35.4242 | -33.6056
23.2145 | -35.4247 | -33.6113
23.2206 | -35.4251 | -33.617
23.2267 | -35.4255 | -33.6228
23.2328 | -35.426 | -33.6284
23.2389 | -35.4265 | -33.6341
23.245 | -35.4269 | -33.6398
23.2511 | -35.4273 | -33.6455
23.2572 | -35.4277 | -33.6512
23.2633 | -35.4282 | -33.6571
23.2694 | -35.4287 | -33.6632
23.2755 | -35.4291 | -33.6686
23.2817 | -35.4297 | -33.6742
23.2878 | -35.4301 | -33.6795
23.2939 | -35.4305 | -33.6857
23.3 | -35.431 | -33.6914
23.3061 | -35.4314 | -33.6965
23.3122 | -35.4318 | -33.7025
23.3183 | -35.4323 | -33.7082
23.3244 | -35.4327 | -33.7141
23.3305 | -35.4332 | -33.7198
23.3366 | -35.4336 | -33.7255
23.3427 | -35.434 | -33.7313
23.3488 | -35.4345 | -33.737
23.3549 | -35.4349 | -33.7428
23.361 | -35.4353 | -33.7484
23.3671 | -35.4358 | -33.7541
23.3732 | -35.4362 | -33.7595
23.3793 | -35.4368 | -33.7654
23.3854 | -35.4372 | -33.7709
23.3915 | -35.4377 | -33.7763
23.3976 | -35.4382 | -33.7818
23.4037 | -35.4387 | -33.7876
23.4098 | -35.4391 | -33.7931
23.4159 | -35.4396 | -33.7991
23.422 | -35.4401 | -33.805
23.4282 | -35.4405 | -33.8104
23.4343 | -35.441 | -33.8163
23.4404 | -35.4414 | -33.8224
23.4465 | -35.4418 | -33.828
23.4526 | -35.4422 | -33.8338
23.4587 | -35.4426 | -33.8396
23.4648 | -35.4431 | -33.845
23.4709 | -35.4435 | -33.851
23.477 | -35.4439 | -33.8568
23.4831 | -35.4443 | -33.8622
23.4892 | -35.4448 | -33.8678
23.4953 | -35.4452 | -33.8739
23.5014 | -35.4457 | -33.8792
23.5075 | -35.4462 | -33.885
23.5136 | -35.4466 | -33.8906
23.5197 | -35.447 | -33.8966
23.5258 | -35.4473 | -33.902
23.5319 | -35.4477 | -33.9078
23.538 | -35.4481 | -33.9136
23.5441 | -35.4485 | -33.9194
23.5502 | -35.449 | -33.9251
23.5563 | -35.4495 | -33.931
23.5624 | -35.4499 | -33.9366
23.5686 | -35.4503 | -33.9426
23.5747 | -35.4508 | -33.9485
23.5808 | -35.4511 | -33.9539
23.5869 | -35.4515 | -33.9595
23.593 | -35.4519 | -33.9655
23.5991 | -35.4524 | -33.9715
23.6052 | -35.4528 | -33.977
23.6113 | -35.4532 | -33.9823
23.6174 | -35.4537 | -33.9877
23.6235 | -35.4542 | -33.9936
23.6296 | -35.4545 | -33.9993
23.6357 | -35.455 | -34.005
23.6418 | -35.4555 | -34.0103
23.6479 | -35.456 | -34.0153
23.654 | -35.4564 | -34.021
23.6601 | -35.4569 | -34.0269
23.6662 | -35.4574 | -34.0326
23.6723 | -35.4579 | -34.0382
23.6784 | -35.4583 | -34.0437
23.6845 | -35.4587 | -34.0489
23.6906 | -35.4591 | -34.0541
23.6967 | -35.4596 | -34.0595
23.7028 | -35.46 | -34.0653
23.7089 | -35.4604 | -34.0711
23.7151 | -35.4609 | -34.0766
23.7212 | -35.4613 | -34.0824
23.7273 | -35.4617 | -34.0886
23.7334 | -35.4622 | -34.094
23.7395 | -35.4626 | -34.0998
23.7456 | -35.4631 | -34.1054
23.7517 | -35.4636 | -34.1108
23.7578 | -35.464 | -34.1168
23.7639 | -35.4644 | -34.1223
23.77 | -35.4649 | -34.128
23.7761 | -35.4654 | -34.1338
23.7822 | -35.4658 | -34.1389
23.7883 | -35.4662 | -34.1445
23.7944 | -35.4667 | -34.1506
23.8005 | -35.4672 | -34.1563
23.8066 | -35.4676 | -34.1617
23.8127 | -35.4679 | -34.1676
23.8188 | -35.4684 | -34.1732
23.8249 | -35.4689 | -34.1791
23.831 | -35.4693 | -34.185
23.8371 | -35.4697 | -34.1914
23.8432 | -35.4702 | -34.197
23.8493 | -35.4707 | -34.2027
23.8555 | -35.4711 | -34.2085
23.8616 | -35.4715 | -34.2143
23.8677 | -35.472 | -34.2197
23.8738 | -35.4724 | -34.2257
23.8799 | -35.4728 | -34.2314
23.886 | -35.4732 | -34.2374
23.8921 | -35.4737 | -34.2433
23.8982 | -35.4741 | -34.2493
23.9043 | -35.4746 | -34.2553
23.9104 | -35.4751 | -34.2611
23.9165 | -35.4756 | -34.2669
23.9226 | -35.4761 | -34.2724
23.9287 | -35.4765 | -34.2783
23.9348 | -35.4769 | -34.284
23.9409 | -35.4774 | -34.2896
23.947 | -35.4779 | -34.2953
23.9531 | -35.4784 | -34.3018
23.9592 | -35.4788 | -34.3078
23.9653 | -35.4794 | -34.3134
23.9714 | -35.4798 | -34.319
23.9775 | -35.4802 | -34.3247
23.9836 | -35.4806 | -34.3306
23.9897 | -35.4812 | -34.3367
23.9958 | -35.4816 | -34.3425
24.002 | -35.482 | -34.3485
24.0081 | -35.4824 | -34.3538
24.0142 | -35.4828 | -34.3594
24.0203 | -35.4833 | -34.365
24.0264 | -35.4837 | -34.3709
24.0325 | -35.4841 | -34.3771
24.0386 | -35.4845 | -34.3828
24.0447 | -35.485 | -34.3888
24.0508 | -35.4854 | -34.3944
24.0569 | -35.4859 | -34.4004
24.063 | -35.4864 | -34.407
24.0691 | -35.4869 | -34.4126
24.0752 | -35.4874 | -34.4181
24.0813 | -35.4879 | -34.4238
24.0874 | -35.4883 | -34.4291
24.0935 | -35.4887 | -34.4347
24.0996 | -35.4891 | -34.4405
24.1057 | -35.4895 | -34.4458
24.1118 | -35.4899 | -34.4513
24.1179 | -35.4903 | -34.4569
24.124 | -35.4907 | -34.4628
24.1301 | -35.4911 | -34.4686
24.1362 | -35.4916 | -34.4743
24.1424 | -35.492 | -34.4799
24.1485 | -35.4925 | -34.4851
24.1546 | -35.4929 | -34.4912
24.1607 | -35.4933 | -34.4972
24.1668 | -35.4937 | -34.503
24.1729 | -35.4942 | -34.5086
24.179 | -35.4946 | -34.5139
24.1851 | -35.4951 | -34.5197
24.1912 | -35.4954 | -34.5254
24.1973 | -35.4959 | -34.5311
24.2034 | -35.4963 | -34.5371
24.2095 | -35.4967 | -34.5426
24.2156 | -35.4971 | -34.5482
24.2217 | -35.4976 | -34.5543
24.2278 | -35.498 | -34.5594
24.2339 | -35.4984 | -34.565
24.24 | -35.4988 | -34.5706
24.2461 | -35.4993 | -34.5766
24.2522 | -35.4996 | -34.5824
24.2583 | -35.5001 | -34.5879
24.2644 | -35.5005 | -34.594
24.2705 | -35.5009 | -34.6001
24.2766 | -35.5013 | -34.6063
24.2827 | -35.5017 | -34.6123
24.2889 | -35.5021 | -34.6183
24.295 | -35.5025 | -34.6241
24.3011 | -35.503 | -34.63
24.3072 | -35.5034 | -34.6355
24.3133 | -35.5038 | -34.6403
24.3194 | -35.5042 | -34.6467
24.3255 | -35.5045 | -34.6527
24.3316 | -35.505 | -34.6586
24.3377 | -35.5054 | -34.6645
24.3438 | -35.5058 | -34.6691
24.3499 | -35.5062 | -34.675
24.356 | -35.5066 | -34.6807
24.3621 | -35.507 | -34.6867
24.3682 | -35.5075 | -34.692
24.3743 | -35.5079 | -34.6974
24.3804 | -35.5083 | -34.7025
24.3865 | -35.5087 | -34.7079
24.3926 | -35.5091 | -34.7135
24.3987 | -35.5096 | -34.7192
24.4048 | -35.51 | -34.725
24.4109 | -35.5104 | -34.7308
24.417 | -35.5108 | -34.7364
24.4231 | -35.5113 | -34.7427
24.4293 | -35.5117 | -34.7481
24.4354 | -35.5122 | -34.7536
24.4415 | -35.5126 | -34.7597
24.4476 | -35.513 | -34.7655
24.4537 | -35.5134 | -34.7709
24.4598 | -35.5138 | -34.7767
24.4659 | -35.5142 | -34.7824
24.472 | -35.5145 | -34.7881
24.4781 | -35.5149 | -34.794
24.4842 | -35.5153 | -34.7997
24.4903 | -35.5158 | -34.8053
24.4964 | -35.5162 | -34.8113
24.5025 | -35.5165 | -34.8168
24.5086 | -35.517 | -34.8227
24.5147 | -35.5174 | -34.8287
24.5208 | -35.5178 | -34.8346
24.5269 | -35.5182 | -34.8403
24.533 | -35.5186 | -34.8463
24.5391 | -35.5191 | -34.8517
24.5452 | -35.5196 | -34.8571
24.5513 | -35.52 | -34.8628
24.5574 | -35.5204 | -34.8683
24.5635 | -35.5208 | -34.8738
24.5696 | -35.5212 | -34.8796
24.5758 | -35.5216 | -34.8852
24.5819 | -35.522 | -34.8906
24.588 | -35.5224 | -34.8973
24.5941 | -35.5229 | -34.903
24.6002 | -35.5233 | -34.9088
24.6063 | -35.5236 | -34.9145
24.6124 | -35.5241 | -34.9208
24.6185 | -35.5245 | -34.9266
24.6246 | -35.5249 | -34.9326
24.6307 | -35.5254 | -34.9385
24.6368 | -35.5258 | -34.9444
24.6429 | -35.5262 | -34.9501
24.649 | -35.5266 | -34.9559
24.6551 | -35.527 | -34.9614
24.6612 | -35.5275 | -34.9677
24.6673 | -35.5279 | -34.9733
24.6734 | -35.5283 | -34.9789
24.6795 | -35.5288 | -34.9847
24.6856 | -35.5292 | -34.9903
24.6917 | -35.5296 | -34.9961
24.6978 | -35.53 | -35.0015
24.7039 | -35.5303 | -35.0077
24.71 | -35.5307 | -35.0137
24.7162 | -35.5312 | -35.0195
24.7223 | -35.5316 | -35.0253
24.7284 | -35.532 | -35.0309
24.7345 | -35.5324 | -35.0368
24.7406 | -35.5329 | -35.0428
24.7467 | -35.5333 | -35.0488
24.7528 | -35.5336 | -35.0542
24.7589 | -35.534 | -35.0597
24.765 | -35.5344 | -35.0651
24.7711 | -35.5348 | -35.0708
24.7772 | -35.5351 | -35.0764
24.7833 | -35.5356 | -35.0822
24.7894 | -35.536 | -35.088
24.7955 | -35.5363 | -35.0935
24.8016 | -35.5367 | -35.0994
24.8077 | -35.5371 | -35.105
24.8138 | -35.5376 | -35.1106
24.8199 | -35.538 | -35.1167
24.826 | -35.5384 | -35.1224
24.8321 | -35.5388 | -35.1281
24.8382 | -35.5393 | -35.1338
24.8443 | -35.5397 | -35.1395
24.8504 | -35.54 | -35.1452
24.8565 | -35.5405 | -35.1512
24.8627 | -35.541 | -35.1567
24.8688 | -35.5414 | -35.1621
24.8749 | -35.5418 | -35.1678
24.881 | -35.5421 | -35.1735
24.8871 | -35.5424 | -35.1794
24.8932 | -35.5429 | -35.1852
24.8993 | -35.5432 | -35.1904
24.9054 | -35.5436 | -35.1965
24.9115 | -35.544 | -35.2023
24.9176 | -35.5444 | -35.2081
24.9237 | -35.5449 | -35.2141
24.9298 | -35.5453 | -35.2199
24.9359 | -35.5457 | -35.2256
24.942 | -35.5461 | -35.2311
24.9481 | -35.5465 | -35.2371
24.9542 | -35.5469 | -35.2429
24.9603 | -35.5473 | -35.2488
24.9664 | -35.5477 | -35.2545
24.9725 | -35.5481 | -35.2602
24.9786 | -35.5485 | -35.2664
24.9847 | -35.5489 | -35.272
24.9908 | -35.5493 | -35.2776
24.9969 | -35.5497 | -35.2831
25.0031 | -35.5501 | -35.2892
25.0092 | -35.5505 | -35.2955
25.0153 | -35.5509 | -35.3013
25.0214 | -35.5514 | -35.3075
25.0275 | -35.5518 | -35.3133
25.0336 | -35.5522 | -35.319
25.0397 | -35.5527 | -35.3249
25.0458 | -35.5531 | -35.3308
25.0519 | -35.5535 | -35.3366
25.058 | -35.5539 | -35.3419
25.0641 | -35.5543 | -35.3478
25.0702 | -35.5547 | -35.3535
25.0763 | -35.5551 | -35.3594
25.0824 | -35.5554 | -35.3652
25.0885 | -35.5559 | -35.3709
25.0946 | -35.5563 | -35.3766
25.1007 | -35.5567 | -35.3821
25.1068 | -35.5571 | -35.3877
25.1129 | -35.5575 | -35.3936
25.119 | -35.5579 | -35.3994
25.1251 | -35.5583 | -35.4052
25.1312 | -35.5587 | -35.4111
25.1373 | -35.5592 | -35.4167
25.1435 | -35.5596 | -35.4224
25.1496 | -35.5599 | -35.4281
25.1557 | -35.5603 | -35.4338
25.1618 | -35.5607 | -35.4395
25.1679 | -35.5611 | -35.4452
25.174 | -35.5615 | -35.451
25.1801 | -35.5619 | -35.4567
25.1862 | -35.5623 | -35.4624
25.1923 | -35.5627 | -35.468
25.1984 | -35.563 | -35.4741
25.2045 | -35.5635 | -35.4802
25.2106 | -35.5639 | -35.4862
25.2167 | -35.5643 | -35.4921
25.2228 | -35.5647 | -35.498
25.2289 | -35.5651 | -35.5036
25.235 | -35.5655 | -35.5095
25.2411 | -35.5659 | -35.5153
25.2472 | -35.5663 | -35.5212
25.2533 | -35.5667 | -35.5269
25.2594 | -35.5671 | -35.5328
25.2655 | -35.5675 | -35.5386
25.2716 | -35.5679 | -35.5445
25.2777 | -35.5683 | -35.5502
25.2838 | -35.5687 | -35.5561
25.29 | -35.569 | -35.5617
25.2961 | -35.5694 | -35.5674
25.3022 | -35.5698 | -35.5729
25.3083 | -35.5702 | -35.5789
25.3144 | -35.5706 | -35.5845
25.3205 | -35.571 | -35.5903
25.3266 | -35.5714 | -35.5958
25.3327 | -35.5718 | -35.6015
25.3388 | -35.5722 | -35.6072
25.3449 | -35.5725 | -35.613
25.351 | -35.5729 | -35.6185
25.3571 | -35.5733 | -35.6244
25.3632 | -35.5736 | -35.6301
25.3693 | -35.5741 | -35.6359
25.3754 | -35.5745 | -35.6416
25.3815 | -35.5748 | -35.6475
25.3876 | -35.5753 | -35.6533
25.3937 | -35.5757 | -35.6593
25.3998 | -35.576 | -35.6649
25.4059 | -35.5764 | -35.6704
25.412 | -35.5767 | -35.6763
25.4181 | -35.5771 | -35.6818
25.4242 | -35.5775 | -35.6879
25.4304 | -35.5779 | -35.6935
25.4365 | -35.5783 | -35.6993
25.4426 | -35.5787 | -35.7049
25.4487 | -35.5791 | -35.7106
25.4548 | -35.5795 | -35.716
25.4609 | -35.5799 | -35.7217
25.467 | -35.5802 | -35.7272
25.4731 | -35.5806 | -35.7331
25.4792 | -35.581 | -35.7391
25.4853 | -35.5815 | -35.745
25.4914 | -35.5819 | -35.7506
25.4975 | -35.5823 | -35.7563
25.5036 | -35.5827 | -35.7621
25.5097 | -35.5831 | -35.7677
25.5158 | -35.5835 | -35.7734
25.5219 | -35.5839 | -35.7794
25.528 | -35.5843 | -35.7852
25.5341 | -35.5847 | -35.791
25.5402 | -35.5851 | -35.7971
25.5463 | -35.5854 | -35.8026
25.5524 | -35.5858 | -35.8086
25.5585 | -35.5861 | -35.814
25.5646 | -35.5866 | -35.8198
25.5707 | -35.5869 | -35.8259
25.5769 | -35.5873 | -35.8316
25.583 | -35.5877 | -35.8371
25.5891 | -35.5881 | -35.8427
25.5952 | -35.5885 | -35.8485
25.6013 | -35.5889 | -35.8542
25.6074 | -35.5892 | -35.8598
25.6135 | -35.5896 | -35.8655
25.6196 | -35.59 | -35.8712
25.6257 | -35.5904 | -35.8772
25.6318 | -35.5908 | -35.8829
25.6379 | -35.5912 | -35.8887
25.644 | -35.5916 | -35.8942
25.6501 | -35.592 | -35.9002
25.6562 | -35.5924 | -35.9059
25.6623 | -35.5928 | -35.9118
25.6684 | -35.5931 | -35.9175
25.6745 | -35.5935 | -35.9234
25.6806 | -35.5939 | -35.929
25.6867 | -35.5943 | -35.9349
25.6928 | -35.5946 | -35.9407
25.6989 | -35.595 | -35.9462
25.705 | -35.5954 | -35.952
25.7111 | -35.5958 | -35.9578
25.7173 | -35.5962 | -35.9631
25.7234 | -35.5966 | -35.9688
25.7295 | -35.597 | -35.9745
25.7356 | -35.5974 | -35.9804
25.7417 | -35.5978 | -35.9861
25.7478 | -35.5982 | -35.9919
25.7539 | -35.5985 | -35.9977
25.76 | -35.5988 | -36.0033
25.7661 | -35.5992 | -36.009
25.7722 | -35.5996 | -36.0146
25.7783 | -35.5999 | -36.0205
25.7844 | -35.6003 | -36.0263
25.7905 | -35.6007 | -36.032
25.7966 | -35.6011 | -36.0378
25.8027 | -35.6014 | -36.0432
25.8088 | -35.6018 | -36.049
25.8149 | -35.6022 | -36.0545
25.821 | -35.6026 | -36.0602
25.8271 | -35.603 | -36.066
25.8332 | -35.6033 | -36.0717
25.8393 | -35.6037 | -36.0773
25.8454 | -35.6041 | -36.0829
25.8515 | -35.6045 | -36.0883
25.8576 | -35.6048 | -36.0939
25.8638 | -35.6052 | -36.0997
25.8699 | -35.6055 | -36.1054
25.876 | -35.6059 | -36.1112
25.8821 | -35.6063 | -36.1169
25.8882 | -35.6066 | -36.1224
25.8943 | -35.607 | -36.1283
25.9004 | -35.6073 | -36.1339
25.9065 | -35.6077 | -36.1395
25.9126 | -35.6081 | -36.1452
25.9187 | -35.6085 | -36.1512
25.9248 | -35.6088 | -36.1563
25.9309 | -35.6092 | -36.162
25.937 | -35.6096 | -36.1675
25.9431 | -35.61 | -36.1732
25.9492 | -35.6103 | -36.1789
25.9553 | -35.6107 | -36.1848
25.9614 | -35.611 | -36.1905
25.9675 | -35.6114 | -36.1963
25.9736 | -35.6118 | -36.2018
25.9797 | -35.6121 | -36.2075
25.9858 | -35.6125 | -36.2133
25.9919 | -35.6129 | -36.2191
25.998 | -35.6132 | -36.2245
26.0042 | -35.6136 | -36.2301
26.0103 | -35.6139 | -36.2357
26.0164 | -35.6143 | -36.2412
26.0225 | -35.6147 | -36.2467
26.0286 | -35.6151 | -36.2523
26.0347 | -35.6155 | -36.258
26.0408 | -35.6158 | -36.2637
26.0469 | -35.6162 | -36.2695
26.053 | -35.6166 | -36.2752
26.0591 | -35.6169 | -36.2808
26.0652 | -35.6173 | -36.2865
26.0713 | -35.6176 | -36.292
26.0774 | -35.618 | -36.2977
26.0835 | -35.6184 | -36.3035
26.0896 | -35.6187 | -36.309
26.0957 | -35.6191 | -36.3145
26.1018 | -35.6195 | -36.3202
26.1079 | -35.6198 | -36.3261
26.114 | -35.6202 | -36.3316
26.1201 | -35.6206 | -36.3374
26.1262 | -35.6209 | -36.3426
26.1323 | -35.6213 | -36.3482
26.1384 | -35.6216 | -36.3538
26.1445 | -35.622 | -36.3594
26.1507 | -35.6224 | -36.3653
26.1568 | -35.6227 | -36.3709
26.1629 | -35.6231 | -36.3767
26.169 | -35.6234 | -36.3823
26.1751 | -35.6237 | -36.3877
26.1812 | -35.6241 | -36.3934
26.1873 | -35.6244 | -36.399
26.1934 | -35.6248 | -36.4047
26.1995 | -35.6252 | -36.4103
26.2056 | -35.6255 | -36.416
26.2117 | -35.6258 | -36.4213
26.2178 | -35.6262 | -36.4268
26.2239 | -35.6266 | -36.4324
26.23 | -35.627 | -36.438
26.2361 | -35.6273 | -36.4438
26.2422 | -35.6276 | -36.4489
26.2483 | -35.628 | -36.4544
26.2544 | -35.6283 | -36.4599
26.2605 | -35.6287 | -36.4655
26.2666 | -35.629 | -36.4714
26.2727 | -35.6294 | -36.4771
26.2788 | -35.6297 | -36.4825
26.2849 | -35.6301 | -36.4882
26.2911 | -35.6305 | -36.4939
26.2972 | -35.6308 | -36.4995
26.3033 | -35.6312 | -36.5053
26.3094 | -35.6316 | -36.511
26.3155 | -35.6319 | -36.5162
26.3216 | -35.6323 | -36.5218
26.3277 | -35.6326 | -36.5277
26.3338 | -35.633 | -36.5335
26.3399 | -35.6333 | -36.539
26.346 | -35.6337 | -36.5444
26.3521 | -35.634 | -36.5502
26.3582 | -35.6343 | -36.5559
26.3643 | -35.6347 | -36.5613
26.3704 | -35.635 | -36.567
26.3765 | -35.6353 | -36.5726
26.3826 | -35.6357 | -36.578
26.3887 | -35.636 | -36.5837
26.3948 | -35.6364 | -36.5894
26.4009 | -35.6368 | -36.595
26.407 | -35.6371 | -36.6006
26.4131 | -35.6375 | -36.6064
26.4192 | -35.6378 | -36.6118
26.4253 | -35.6382 | -36.6173
26.4314 | -35.6385 | -36.6229
26.4376 | -35.6389 | -36.6284
26.4437 | -35.6392 | -36.634
26.4498 | -35.6396 | -36.6398
26.4559 | -35.6399 | -36.6453
26.462 | -35.6402 | -36.6509
26.4681 | -35.6405 | -36.6564
26.4742 | -35.6409 | -36.6617
26.4803 | -35.6412 | -36.6673
26.4864 | -35.6416 | -36.6727
26.4925 | -35.6419 | -36.6783
26.4986 | -35.6423 | -36.6839
26.5047 | -35.6427 | -36.6895
26.5108 | -35.643 | -36.6951
26.5169 | -35.6433 | -36.7006
26.523 | -35.6436 | -36.7061
26.5291 | -35.644 | -36.7116
26.5352 | -35.6443 | -36.717
26.5413 | -35.6446 | -36.7228
26.5474 | -35.645 | -36.7281
26.5535 | -35.6454 | -36.7338
26.5596 | -35.6457 | -36.7392
26.5657 | -35.646 | -36.7448
26.5718 | -35.6464 | -36.7502
26.578 | -35.6468 | -36.7558
26.5841 | -35.6471 | -36.7616
26.5902 | -35.6474 | -36.7671
26.5963 | -35.6477 | -36.7725
26.6024 | -35.6481 | -36.7781
26.6085 | -35.6484 | -36.7835
26.6146 | -35.6488 | -36.7895
26.6207 | -35.6492 | -36.795
26.6268 | -35.6495 | -36.8008
26.6329 | -35.6498 | -36.8061
26.639 | -35.6501 | -36.8121
26.6451 | -35.6504 | -36.8178
26.6512 | -35.6508 | -36.8232
26.6573 | -35.6511 | -36.8286
26.6634 | -35.6515 | -36.8342
26.6695 | -35.6518 | -36.8398
26.6756 | -35.6522 | -36.8453
26.6817 | -35.6525 | -36.851
26.6878 | -35.6528 | -36.8564
26.6939 | -35.6531 | -36.8618
26.7 | -35.6535 | -36.867
26.7061 | -35.6538 | -36.8727
26.7122 | -35.6541 | -36.8782
26.7183 | -35.6545 | -36.8837
26.7245 | -35.6548 | -36.8894
26.7306 | -35.6551 | -36.8951
26.7367 | -35.6554 | -36.9005
26.7428 | -35.6558 | -36.906
26.7489 | -35.6561 | -36.9117
26.755 | -35.6564 | -36.9172
26.7611 | -35.6568 | -36.923
26.7672 | -35.6571 | -36.9286
26.7733 | -35.6574 | -36.934
26.7794 | -35.6577 | -36.9396
26.7855 | -35.6581 | -36.9451
26.7916 | -35.6584 | -36.9505
26.7977 | -35.6587 | -36.956
26.8038 | -35.6591 | -36.9614
26.8099 | -35.6594 | -36.9668
26.816 | -35.6597 | -36.9725
26.8221 | -35.66 | -36.978
26.8282 | -35.6604 | -36.9836
26.8343 | -35.6607 | -36.9893
26.8404 | -35.661 | -36.9947
26.8465 | -35.6614 | -37.0003
26.8526 | -35.6617 | -37.0057
26.8587 | -35.662 | -37.0111
26.8649 | -35.6623 | -37.0168
26.871 | -35.6627 | -37.0225
26.8771 | -35.663 | -37.0279
26.8832 | -35.6633 | -37.0332
26.8893 | -35.6637 | -37.0386
26.8954 | -35.664 | -37.0442
26.9015 | -35.6643 | -37.0498
26.9076 | -35.6646 | -37.0554
26.9137 | -35.665 | -37.0607
26.9198 | -35.6653 | -37.0663
26.9259 | -35.6656 | -37.0719
26.932 | -35.6659 | -37.0774
26.9381 | -35.6662 | -37.0828
26.9442 | -35.6665 | -37.0881
26.9503 | -35.6668 | -37.0937
26.9564 | -35.6671 | -37.0991
26.9625 | -35.6674 | -37.1047
26.9686 | -35.6678 | -37.1102
26.9747 | -35.6681 | -37.1158
26.9808 | -35.6684 | -37.1211
26.9869 | -35.6687 | -37.1263
26.993 | -35.669 | -37.132
26.9991 | -35.6693 | -37.1377
27.0052 | -35.6696 | -37.1432
27.0114 | -35.6699 | -37.1487
27.0175 | -35.6702 | -37.1539
27.0236 | -35.6705 | -37.1593
27.0297 | -35.6708 | -37.1648
27.0358 | -35.6712 | -37.1702
27.0419 | -35.6715 | -37.1758
27.048 | -35.6718 | -37.1812
27.0541 | -35.6721 | -37.1864
27.0602 | -35.6725 | -37.1918
27.0663 | -35.6728 | -37.1974
27.0724 | -35.6731 | -37.2028
27.0785 | -35.6734 | -37.2085
27.0846 | -35.6737 | -37.2134
27.0907 | -35.674 | -37.2188
27.0968 | -35.6743 | -37.2242
27.1029 | -35.6746 | -37.2297
27.109 | -35.6749 | -37.2351
27.1151 | -35.6752 | -37.2404
27.1212 | -35.6755 | -37.2456
27.1273 | -35.6758 | -37.2511
27.1334 | -35.6761 | -37.2566
27.1395 | -35.6764 | -37.262
27.1456 | -35.6767 | -37.2674
27.1518 | -35.6771 | -37.2729
27.1579 | -35.6774 | -37.2783
27.164 | -35.6777 | -37.284
27.1701 | -35.678 | -37.2893
27.1762 | -35.6783 | -37.2948
27.1823 | -35.6786 | -37.3002
27.1884 | -35.6789 | -37.3059
27.1945 | -35.6792 | -37.3113
27.2006 | -35.6796 | -37.3166
27.2067 | -35.6799 | -37.3222
27.2128 | -35.6802 | -37.3273
27.2189 | -35.6805 | -37.3327
27.225 | -35.6809 | -37.3383
27.2311 | -35.6811 | -37.3435
27.2372 | -35.6815 | -37.3488
27.2433 | -35.6818 | -37.354
27.2494 | -35.6821 | -37.3596
27.2555 | -35.6824 | -37.365
27.2616 | -35.6827 | -37.3701
27.2677 | -35.683 | -37.3757
27.2738 | -35.6833 | -37.3812
27.2799 | -35.6836 | -37.3867
27.286 | -35.6839 | -37.392
27.2921 | -35.6842 | -37.3976
27.2983 | -35.6845 | -37.403
27.3044 | -35.6848 | -37.4084
27.3105 | -35.6851 | -37.414
27.3166 | -35.6854 | -37.4193
27.3227 | -35.6857 | -37.4247
27.3288 | -35.6861 | -37.4299
27.3349 | -35.6863 | -37.4349
27.341 | -35.6866 | -37.4405
27.3471 | -35.687 | -37.4459
27.3532 | -35.6873 | -37.4514
27.3593 | -35.6876 | -37.4567
27.3654 | -35.688 | -37.4621
27.3715 | -35.6883 | -37.4671
27.3776 | -35.6886 | -37.4725
27.3837 | -35.689 | -37.478
27.3898 | -35.6893 | -37.4833
27.3959 | -35.6896 | -37.4889
27.402 | -35.6899 | -37.4941
27.4081 | -35.6902 | -37.4995
27.4142 | -35.6905 | -37.5049
27.4203 | -35.6908 | -37.5102
27.4264 | -35.6911 | -37.5153
27.4325 | -35.6914 | -37.5208
27.4387 | -35.6917 | -37.526
27.4448 | -35.6919 | -37.5314
27.4509 | -35.6923 | -37.537
27.457 | -35.6926 | -37.5424
27.4631 | -35.6929 | -37.548
27.4692 | -35.6932 | -37.5533
27.4753 | -35.6935 | -37.5585
27.4814 | -35.6938 | -37.5639
27.4875 | -35.6941 | -37.5694
27.4936 | -35.6944 | -37.5747
27.4997 | -35.6947 | -37.58
27.5058 | -35.695 | -37.5855
27.5119 | -35.6953 | -37.5904
27.518 | -35.6956 | -37.596
27.5241 | -35.6959 | -37.6013
27.5302 | -35.6962 | -37.607
27.5363 | -35.6964 | -37.6122
27.5424 | -35.6967 | -37.6173
27.5485 | -35.697 | -37.6225
27.5546 | -35.6973 | -37.6278
27.5607 | -35.6976 | -37.6333
27.5668 | -35.6979 | -37.6385
27.5729 | -35.6982 | -37.6438
27.5791 | -35.6985 | -37.6489
27.5852 | -35.6988 | -37.6543
27.5913 | -35.6991 | -37.6593
27.5974 | -35.6994 | -37.6646
27.6035 | -35.6997 | -37.6702
27.6096 | -35.6999 | -37.6753
27.6157 | -35.7002 | -37.6804
27.6218 | -35.7005 | -37.6861
27.6279 | -35.7007 | -37.6914
27.634 | -35.701 | -37.6969
27.6401 | -35.7013 | -37.7022
27.6462 | -35.7017 | -37.7077
27.6523 | -35.7019 | -37.7129
27.6584 | -35.7022 | -37.7183
27.6645 | -35.7025 | -37.7236
27.6706 | -35.7028 | -37.729
27.6767 | -35.7032 | -37.7345
27.6828 | -35.7034 | -37.7396
27.6889 | -35.7038 | -37.7446
27.695 | -35.7041 | -37.7496
27.7011 | -35.7044 | -37.755
27.7072 | -35.7047 | -37.7604
27.7133 | -35.7049 | -37.7656
27.7194 | -35.7052 | -37.771
27.7256 | -35.7055 | -37.7762
27.7317 | -35.7058 | -37.7815
27.7378 | -35.7061 | -37.7867
27.7439 | -35.7064 | -37.7922
27.75 | -35.7067 | -37.7977
27.7561 | -35.707 | -37.8028
27.7622 | -35.7073 | -37.8082
27.7683 | -35.7076 | -37.8137
27.7744 | -35.7079 | -37.8191
27.7805 | -35.7082 | -37.8242
27.7866 | -35.7085 | -37.8295
27.7927 | -35.7087 | -37.8346
27.7988 | -35.709 | -37.8399
27.8049 | -35.7093 | -37.8453
27.811 | -35.7096 | -37.8505
27.8171 | -35.7099 | -37.8562
27.8232 | -35.7102 | -37.861
27.8293 | -35.7104 | -37.8661
27.8354 | -35.7107 | -37.8716
27.8415 | -35.711 | -37.8769
27.8476 | -35.7113 | -37.8822
27.8537 | -35.7117 | -37.8875
27.8598 | -35.7119 | -37.8927
27.866 | -35.7122 | -37.8978
27.8721 | -35.7125 | -37.903
27.8782 | -35.7128 | -37.9084
27.8843 | -35.7131 | -37.9141
27.8904 | -35.7133 | -37.9194
27.8965 | -35.7136 | -37.9245
27.9026 | -35.7139 | -37.93
27.9087 | -35.7142 | -37.9352
27.9148 | -35.7145 | -37.9404
27.9209 | -35.7147 | -37.9458
27.927 | -35.715 | -37.9513
27.9331 | -35.7153 | -37.9564
27.9392 | -35.7155 | -37.9619
27.9453 | -35.7158 | -37.9672
27.9514 | -35.7161 | -37.9724
27.9575 | -35.7164 | -37.978
27.9636 | -35.7167 | -37.9832
27.9697 | -35.717 | -37.9885
27.9758 | -35.7172 | -37.9938
27.9819 | -35.7175 | -37.9993
27.988 | -35.7178 | -38.0043
27.9941 | -35.7181 | -38.0099
28.0002 | -35.7184 | -38.015
28.0063 | -35.7186 | -38.0204
28.0125 | -35.7189 | -38.0257
28.0186 | -35.7192 | -38.0308
28.0247 | -35.7195 | -38.036
28.0308 | -35.7198 | -38.0411
28.0369 | -35.72 | -38.046
28.043 | -35.7203 | -38.0514
28.0491 | -35.7206 | -38.0566
28.0552 | -35.7209 | -38.0618
28.0613 | -35.7211 | -38.0666
28.0674 | -35.7214 | -38.0718
28.0735 | -35.7217 | -38.0768

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.2985 | -42.1553
21.4908 | -35.4072 | -35.1906
16.4948 | -34.7944 | -27.7609
12.6602 | -34.4378 | -21.4581
9.7171 | -34.2221 | -16.5338
7.4582 | -34.0734 | -12.5508
5.7244 | -33.9704 | -9.3317
4.3936 | -33.9082 | -6.7209
3.3722 | -33.8948 | -4.6731
2.5883 | -33.8945 | -3.1184
1.9866 | -33.8937 | -1.8803
1.5248 | -33.9016 | -0.88738
1.1703 | -33.9054 | -0.23652
1 | -33.9088 | 0.085792
0.89822 | -33.8996 | 0.35596
0.68941 | -33.8931 | 0.93449
0.52919 | -33.9023 | 1.3837
0.40612 | -33.9002 | 1.9901
0.31171 | -33.9016 | 2.4388
0.23928 | -33.9005 | 3.1654
0.18358 | -33.9231 | 3.5852
0.14091 | -33.9365 | 4.4764
0.10819 | -33.9663 | 5.6318
0.083017 | -34.0082 | 7.1046
0.063705 | -34.0397 | 8.9899
0.048923 | -34.1184 | 11.4858
0.037479 | -34.2388 | 14.6922
0.028801 | -34.4276 | 18.7217
0.022078 | -34.724 | 24.0816
0.017023 | -35.2063 | 30.5109
0.013018 | -36.0407 | 38.214
0.010014 | -37.1483 | 46.1184

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003658 | -13.6253
0.007316 | -34.282
0.010974 | -39.0025
0.014633 | -41.0447
0.018291 | -44.0134
0.021949 | -45.6754
0.025607 | -47.7561
0.029265 | -50.225
0.032923 | -50.3231
0.036581 | -52.0707
0.04024 | -54.4678
0.043898 | -55.4482
0.047556 | -56.7844
0.051214 | -57.7654
0.054872 | -57.0765
0.05853 | -56.9324
0.062188 | -58.0504
0.065847 | -58.9966
0.069505 | -58.9382
0.073163 | -59.5224
0.076821 | -59.153
0.080479 | -59.6028
0.084137 | -60.7357
0.087796 | -62.3873
0.091454 | -62.9692
0.095112 | -63.3218
0.09877 | -63.4822
0.10243 | -63.4377
0.10609 | -64.5059
0.10974 | -65.5053
0.1134 | -66.6892
0.11706 | -65.6237
0.12072 | -65.4051
0.12438 | -68.2279
0.12804 | -67.5408
0.13169 | -67.6623
0.13535 | -67.3158
0.13901 | -67.1938
0.14267 | -67.8644
0.14633 | -68.3274
0.14998 | -67.4185
0.15364 | -67.671
0.1573 | -69.0504
0.16096 | -69.5437
0.16462 | -70.4078
0.16828 | -70.341
0.17193 | -70.6585
0.17559 | -70.3536
0.17925 | -70.5212
0.18291 | -68.6063
0.18657 | -68.2195
0.19022 | -69.5681
0.19388 | -70.5168
0.19754 | -70.651
0.2012 | -71.8835
0.20486 | -72.3194
0.20851 | -71.8735
0.21217 | -72.6828
0.21583 | -73.4484
0.21949 | -72.9673
0.22315 | -73.5067
0.22681 | -73.0736
0.23046 | -73.5474
0.23412 | -72.3107
0.23778 | -72.554
0.24144 | -72.713
0.2451 | -72.9502
0.24875 | -72.4712
0.25241 | -73.9275
0.25607 | -74.0135
0.25973 | -74.6953
0.26339 | -74.6984
0.26704 | -73.9188
0.2707 | -74.1094
0.27436 | -74.5814
0.27802 | -74.0683
0.28168 | -74.3212
0.28534 | -75.6419
0.28899 | -75.9762
0.29265 | -75.9042
0.29631 | -76.1641
0.29997 | -75.9697
0.30363 | -75.6663
0.30728 | -74.9811
0.31094 | -74.1292
0.3146 | -75.626
0.31826 | -76.1239
0.32192 | -75.9969
0.32558 | -76.2456
0.32923 | -76.4984
0.33289 | -76.4561
0.33655 | -76.9846
0.34021 | -76.9132
0.34387 | -76.3216
0.34752 | -76.6718
0.35118 | -77.3164
0.35484 | -77.2509
0.3585 | -76.5967
0.36216 | -76.6714
0.36582 | -76.8547
0.36947 | -77.6999
0.37313 | -77.4431
0.37679 | -76.8428
0.38045 | -78.3059
0.3841 | -78.3158
0.38776 | -77.787
0.39142 | -78.5513
0.39508 | -78.0526
0.39874 | -77.7221
0.4024 | -77.4667
0.40605 | -78.0404
0.40971 | -78.3308
0.41337 | -78.9169
0.41703 | -79.2513
0.42069 | -79.7713
0.42435 | -79.7173
0.428 | -79.3255
0.43166 | -79.2652
0.43532 | -78.6287
0.43898 | -79.0951
0.44264 | -79.1485
0.44629 | -79.5325
0.44995 | -79.8022
0.45361 | -80.0246
0.45727 | -80.0057
0.46093 | -79.881
0.46459 | -79.67
0.46824 | -79.8057
0.4719 | -78.9712
0.47556 | -80.0165
0.47922 | -80.313
0.48287 | -80.0973
0.48653 | -79.3962
0.49019 | -78.916
0.49385 | -78.3634
0.49751 | -79.8763
0.50117 | -82.1533
0.50482 | -82.3852
0.50848 | -82.2125
0.51214 | -80.5812
0.5158 | -81.0859
0.51946 | -81.4876
0.52311 | -79.9937
0.52677 | -80.2116
0.53043 | -80.4464
0.53409 | -80.5813
0.53775 | -80.7392
0.54141 | -80.9676
0.54506 | -81.3731
0.54872 | -81.8344
0.55238 | -81.9665
0.55604 | -81.2994
0.5597 | -80.6986
0.56336 | -80.9857
0.56701 | -81.3844
0.57067 | -81.1322
0.57433 | -81.2521
0.57799 | -81.9616
0.58164 | -82.471
0.5853 | -82.5183
0.58896 | -82.5825
0.59262 | -82.5876
0.59628 | -82.4954
0.59994 | -82.451
0.60359 | -82.493
0.60725 | -82.3999
0.61091 | -82.3293
0.61457 | -82.4573
0.61823 | -82.4843
0.62189 | -82.1076
0.62554 | -82.0202
0.6292 | -82.255
0.63286 | -82.5668
0.63652 | -82.7948
0.64018 | -82.7458
0.64383 | -82.8205
0.64749 | -82.723
0.65115 | -82.5521
0.65481 | -82.7448
0.65847 | -83.2066
0.66212 | -83.5708
0.66578 | -83.7176
0.66944 | -83.8756
0.6731 | -84.3431
0.67676 | -84.3806
0.68041 | -83.6173
0.68407 | -82.8763
0.68773 | -82.7063
0.69139 | -83.0513
0.69505 | -83.505
0.69871 | -83.9336
0.70236 | -84.1844
0.70602 | -84.1077
0.70968 | -83.9345
0.71334 | -83.919
0.717 | -83.8007
0.72066 | -84.0017
0.72431 | -84.6117
0.72797 | -84.864
0.73163 | -84.6678
0.73529 | -84.7014
0.73895 | -85.0211
0.7426 | -85.2876
0.74626 | -84.9272
0.74992 | -83.9293
0.75358 | -83.3167
0.75724 | -83.4511
0.76089 | -83.9929
0.76455 | -84.221
0.76821 | -84.0823
0.77187 | -84.4556
0.77553 | -85.0364
0.77919 | -84.9445
0.78284 | -84.4398
0.7865 | -84.3938
0.79016 | -84.6038
0.79382 | -84.7981
0.79748 | -85.2336
0.80113 | -85.5741
0.80479 | -85.1834
0.80845 | -84.9034
0.81211 | -85.3636
0.81577 | -85.9256
0.81942 | -85.9342
0.82308 | -85.3397
0.82674 | -84.9263
0.8304 | -85.0885
0.83406 | -84.7888
0.83771 | -84.3486
0.84137 | -84.5978
0.84503 | -85.2852
0.84869 | -85.5095
0.85235 | -85.2688
0.85601 | -84.8762
0.85966 | -84.5462
0.86332 | -84.702
0.86698 | -85.2298
0.87064 | -85.5906
0.8743 | -85.6264
0.87796 | -85.531
0.88161 | -85.4807
0.88527 | -85.6743
0.88893 | -86.2785
0.89259 | -86.6771
0.89625 | -86.673
0.8999 | -86.242
0.90356 | -85.898
0.90722 | -86.4949
0.91088 | -87.4816
0.91454 | -87.5108
0.91819 | -86.8853
0.92185 | -86.5745
0.92551 | -87.0712
0.92917 | -87.0922
0.93283 | -86.3128
0.93649 | -85.7765
0.94014 | -85.7765
0.9438 | -86.1019
0.94746 | -85.9595
0.95112 | -85.4661
0.95478 | -85.4903
0.95843 | -86.0827
0.96209 | -86.8666
0.96575 | -87.2955
0.96941 | -87.4811
0.97307 | -87.4469
0.97672 | -87.1689
0.98038 | -87.1354
0.98404 | -87.3546
0.9877 | -87.6071
0.99136 | -87.9317
0.99502 | -87.9511
0.99867 | -87.6079
1.0023 | -87.3116
1.006 | -87.5808
1.0096 | -87.7918
1.0133 | -87.6406
1.017 | -87.3359
1.0206 | -87.0637
1.0243 | -87.0839
1.0279 | -87.1812
1.0316 | -87.5386
1.0353 | -87.6466
1.0389 | -87.7113
1.0426 | -87.7884
1.0462 | -87.6456
1.0499 | -87.4221
1.0535 | -87.4869
1.0572 | -87.7221
1.0609 | -87.9292
1.0645 | -87.8408
1.0682 | -87.7228
1.0718 | -87.8553
1.0755 | -88.0098
1.0792 | -87.9264
1.0828 | -87.7746
1.0865 | -87.7765
1.0901 | -87.9152
1.0938 | -88.0098
1.0974 | -87.896
1.1011 | -87.7896
1.1048 | -87.9378
1.1084 | -88.1499
1.1121 | -88.1124
1.1157 | -87.925
1.1194 | -87.8869
1.1231 | -88.0557
1.1267 | -88.2235
1.1304 | -88.2822
1.134 | -88.2939
1.1377 | -88.3315
1.1413 | -88.459
1.145 | -88.6201
1.1487 | -88.7126
1.1523 | -88.6651
1.156 | -88.5243
1.1596 | -88.4042
1.1633 | -88.3105
1.1669 | -88.3057
1.1706 | -88.4968
1.1743 | -88.7412
1.1779 | -88.8722
1.1816 | -89.0533
1.1852 | -89.4707
1.1889 | -89.9182
1.1926 | -89.9885
1.1962 | -89.5701
1.1999 | -89.0334
1.2035 | -88.8106
1.2072 | -88.9523
1.2108 | -89.2367
1.2145 | -89.4474
1.2182 | -89.4614
1.2218 | -89.3239
1.2255 | -89.1485
1.2291 | -88.9704
1.2328 | -88.8454
1.2365 | -88.7854
1.2401 | -88.8123
1.2438 | -88.9641
1.2474 | -89.1819
1.2511 | -89.4197
1.2547 | -89.6283
1.2584 | -89.6217
1.2621 | -89.2872
1.2657 | -88.8742
1.2694 | -88.774
1.273 | -88.977
1.2767 | -89.0773
1.2804 | -88.9259
1.284 | -88.7675
1.2877 | -88.7595
1.2913 | -88.8044
1.295 | -88.7986
1.2986 | -88.8527
1.3023 | -89.1066
1.306 | -89.4743
1.3096 | -89.597
1.3133 | -89.314
1.3169 | -88.96
1.3206 | -88.7973
1.3242 | -88.9048
1.3279 | -89.2037
1.3316 | -89.457
1.3352 | -89.6771
1.3389 | -90.0253
1.3425 | -90.3605
1.3462 | -90.3601
1.3499 | -89.9598
1.3535 | -89.4815
1.3572 | -89.3415
1.3608 | -89.611
1.3645 | -89.9234
1.3681 | -89.9783
1.3718 | -89.8579
1.3755 | -89.7128
1.3791 | -89.6852
1.3828 | -89.8252
1.3864 | -90.0711
1.3901 | -90.2447
1.3938 | -90.178
1.3974 | -90.0005
1.4011 | -89.9767
1.4047 | -90.1468
1.4084 | -90.3474
1.412 | -90.5064
1.4157 | -90.7174
1.4194 | -90.9695
1.423 | -91.0763
1.4267 | -90.8974
1.4303 | -90.5957
1.434 | -90.4495
1.4377 | -90.3792
1.4413 | -90.043
1.445 | -89.5387
1.4486 | -89.2724
1.4523 | -89.37
1.4559 | -89.6687
1.4596 | -89.9357
1.4633 | -90.0551
1.4669 | -90.0098
1.4706 | -89.9133
1.4742 | -89.952
1.4779 | -90.0733
1.4815 | -90.1373
1.4852 | -90.2248
1.4889 | -90.4051
1.4925 | -90.6219
1.4962 | -90.8202
1.4998 | -90.9387
1.5035 | -90.8986
1.5072 | -90.7337
1.5108 | -90.5386
1.5145 | -90.3055
1.5181 | -90.0779
1.5218 | -90.0781
1.5254 | -90.4052
1.5291 | -90.8499
1.5328 | -91.0545
1.5364 | -90.9477
1.5401 | -90.7454
1.5437 | -90.5703
1.5474 | -90.4348
1.5511 | -90.3743
1.5547 | -90.4511
1.5584 | -90.6737
1.562 | -90.9642
1.5657 | -91.2133
1.5693 | -91.3341
1.573 | -91.4679
1.5767 | -91.6562
1.5803 | -91.5711
1.584 | -91.45
1.5876 | -91.262
1.5913 | -91.0528
1.595 | -90.8627
1.5986 | -90.8089
1.6023 | -90.9339
1.6059 | -91.167
1.6096 | -91.3856
1.6132 | -91.5204
1.6169 | -91.5743
1.6206 | -91.5772
1.6242 | -91.5643
1.6279 | -91.596
1.6315 | -91.7227
1.6352 | -91.9774
1.6388 | -92.274
1.6425 | -92.4455
1.6462 | -92.4047
1.6498 | -92.2205
1.6535 | -92.0397
1.6571 | -91.9713
1.6608 | -92.0043
1.6645 | -92.0427
1.6681 | -92.0028
1.6718 | -91.8847
1.6754 | -91.7528
1.6791 | -91.6717
1.6827 | -91.6321
1.6864 | -91.5474
1.6901 | -91.3999
1.6937 | -91.3088
1.6974 | -91.3661
1.701 | -91.5327
1.7047 | -91.7691
1.7084 | -92.099
1.712 | -92.5206
1.7157 | -92.8843
1.7193 | -93.0172
1.723 | -92.8382
1.7266 | -92.4446
1.7303 | -92.021
1.734 | -91.7569
1.7376 | -91.7648
1.7413 | -92.0443
1.7449 | -92.4684
1.7486 | -92.8572
1.7523 | -93.031
1.7559 | -92.9223
1.7596 | -92.5793
1.7632 | -92.2028
1.7669 | -91.9831
1.7705 | -91.97
1.7742 | -92.1201
1.7779 | -92.3958
1.7815 | -92.7211
1.7852 | -92.9605
1.7888 | -92.9736
1.7925 | -92.7524
1.7961 | -92.3757
1.7998 | -92.0162
1.8035 | -91.8564
1.8071 | -91.9152
1.8108 | -92.0685
1.8144 | -92.1817
1.8181 | -92.2174
1.8218 | -92.2368
1.8254 | -92.3139
1.8291 | -92.455
1.8327 | -92.5003
1.8364 | -92.292
1.84 | -91.9083
1.8437 | -91.6125
1.8474 | -91.6185
1.851 | -91.943
1.8547 | -92.3836
1.8583 | -92.6298
1.862 | -92.5347
1.8657 | -92.2874
1.8693 | -92.1449
1.873 | -92.1828
1.8766 | -92.2765
1.8803 | -92.2995
1.8839 | -92.2435
1.8876 | -92.1559
1.8913 | -92.0675
1.8949 | -92.0642
1.8986 | -92.222
1.9022 | -92.5069
1.9059 | -92.8253
1.9096 | -93.1164
1.9132 | -93.314
1.9169 | -93.3721
1.9205 | -93.3156
1.9242 | -93.1832
1.9278 | -92.9578
1.9315 | -92.6744
1.9352 | -92.4722
1.9388 | -92.4524
1.9425 | -92.5673
1.9461 | -92.6809
1.9498 | -92.7123
1.9534 | -92.6754
1.9571 | -92.6112
1.9608 | -92.5429
1.9644 | -92.5192
1.9681 | -92.6062
1.9717 | -92.8029
1.9754 | -93.0407
1.9791 | -93.2554
1.9827 | -93.4352
1.9864 | -93.576
1.99 | -93.6079
1.9937 | -93.4842
1.9973 | -93.2719
2.001 | -93.1083
2.0047 | -93.0807
2.0083 | -93.1583
2.012 | -93.2609
2.0156 | -93.3649
2.0193 | -93.5193
2.023 | -93.7388
2.0266 | -93.9105
2.0303 | -93.87
2.0339 | -93.6454
2.0376 | -93.4584
2.0412 | -93.4704
2.0449 | -93.6898
2.0486 | -93.9981
2.0522 | -94.2528
2.0559 | -94.346
2.0595 | -94.2765
2.0632 | -94.1222
2.0669 | -93.9157
2.0705 | -93.6899
2.0742 | -93.5174
2.0778 | -93.4326
2.0815 | -93.4345
2.0851 | -93.4985
2.0888 | -93.6071
2.0925 | -93.7121
2.0961 | -93.7701
2.0998 | -93.8027
2.1034 | -93.8641
2.1071 | -93.9685
2.1107 | -94.0614
2.1144 | -94.1122
2.1181 | -94.0971
2.1217 | -94.0081
2.1254 | -93.854
2.129 | -93.6626
2.1327 | -93.495
2.1364 | -93.3992
2.14 | -93.3986
2.1437 | -93.4839
2.1473 | -93.5694
2.151 | -93.5874
2.1546 | -93.5228
2.1583 | -93.4175
2.162 | -93.3222
2.1656 | -93.2924
2.1693 | -93.3578
2.1729 | -93.5284
2.1766 | -93.7952
2.1803 | -94.1002
2.1839 | -94.3414
2.1876 | -94.4735
2.1912 | -94.4992
2.1949 | -94.4949
2.1985 | -94.5104
2.2022 | -94.5578
2.2059 | -94.6076
2.2095 | -94.6257
2.2132 | -94.5995
2.2168 | -94.5261
2.2205 | -94.4314
2.2242 | -94.3463
2.2278 | -94.2737
2.2315 | -94.214
2.2351 | -94.1763
2.2388 | -94.1508
2.2424 | -94.1055
2.2461 | -94.0217
2.2498 | -93.9203
2.2534 | -93.8056
2.2571 | -93.6861
2.2607 | -93.5678
2.2644 | -93.4655
2.2681 | -93.398
2.2717 | -93.4025
2.2754 | -93.5317
2.279 | -93.828
2.2827 | -94.2658
2.2863 | -94.7284
2.29 | -95.0243
2.2937 | -95.0333
2.2973 | -94.7838
2.301 | -94.3937
2.3046 | -94.0089
2.3083 | -93.72
2.3119 | -93.5684
2.3156 | -93.5606
2.3193 | -93.6772
2.3229 | -93.8704
2.3266 | -94.095
2.3302 | -94.3157
2.3339 | -94.4762
2.3376 | -94.53
2.3412 | -94.4963
2.3449 | -94.4191
2.3485 | -94.3522
2.3522 | -94.3148
2.3558 | -94.2906
2.3595 | -94.2582
2.3632 | -94.2218
2.3668 | -94.2009
2.3705 | -94.225
2.3741 | -94.3359
2.3778 | -94.531
2.3815 | -94.7985
2.3851 | -95.0862
2.3888 | -95.3104
2.3924 | -95.4042
2.3961 | -95.3293
2.3997 | -95.1103
2.4034 | -94.84
2.4071 | -94.6026
2.4107 | -94.4296
2.4144 | -94.3248
2.418 | -94.3251
2.4217 | -94.406
2.4254 | -94.5296
2.429 | -94.5977
2.4327 | -94.5603
2.4363 | -94.4615
2.44 | -94.3642
2.4436 | -94.3125
2.4473 | -94.3304
2.451 | -94.4167
2.4546 | -94.5357
2.4583 | -94.629
2.4619 | -94.6595
2.4656 | -94.5814
2.4692 | -94.4054
2.4729 | -94.1694
2.4766 | -93.952
2.4802 | -93.8374
2.4839 | -93.8792
2.4875 | -94.0698
2.4912 | -94.3217
2.4949 | -94.5157
2.4985 | -94.5825
2.5022 | -94.552
2.5058 | -94.52
2.5095 | -94.5461
2.5131 | -94.6289
2.5168 | -94.7198
2.5205 | -94.767
2.5241 | -94.7528
2.5278 | -94.689
2.5314 | -94.6038
2.5351 | -94.5399
2.5388 | -94.5491
2.5424 | -94.656
2.5461 | -94.8166
2.5497 | -94.9576
2.5534 | -95.0514
2.557 | -95.1427
2.5607 | -95.2791
2.5644 | -95.4591
2.568 | -95.6354
2.5717 | -95.7378
2.5753 | -95.7045
2.579 | -95.5359
2.5827 | -95.3052
2.5863 | -95.1269
2.59 | -95.0864
2.5936 | -95.1504
2.5973 | -95.2101
2.6009 | -95.2026
2.6046 | -95.1582
2.6083 | -95.148
2.6119 | -95.2232
2.6156 | -95.3749
2.6192 | -95.5624
2.6229 | -95.7413
2.6265 | -95.8552
2.6302 | -95.8571
2.6339 | -95.7492
2.6375 | -95.5416
2.6412 | -95.4006
2.6448 | -95.2871
2.6485 | -95.2312
2.6522 | -95.2392
2.6558 | -95.2891
2.6595 | -95.3685
2.6631 | -95.445
2.6668 | -95.5167
2.6704 | -95.5935
2.6741 | -95.6739
2.6778 | -95.7587
2.6814 | -95.8542
2.6851 | -95.9501
2.6887 | -96.031
2.6924 | -96.0907
2.6961 | -96.1067
2.6997 | -96.0563
2.7034 | -95.925
2.707 | -95.7507
2.7107 | -95.578
2.7143 | -95.4299
2.718 | -95.3463
2.7217 | -95.3483
2.7253 | -95.4148
2.729 | -95.5246
2.7326 | -95.6529
2.7363 | -95.7839
2.74 | -95.9012
2.7436 | -96.0017
2.7473 | -96.0826
2.7509 | -96.1335
2.7546 | -96.1472
2.7582 | -96.139
2.7619 | -96.1298
2.7656 | -96.1165
2.7692 | -96.0909
2.7729 | -96.0621
2.7765 | -96.0208
2.7802 | -95.9611
2.7838 | -95.8673
2.7875 | -95.7425
2.7912 | -95.6084
2.7948 | -95.4965
2.7985 | -95.4413
2.8021 | -95.4516
2.8058 | -95.5131
2.8095 | -95.6018
2.8131 | -95.6986
2.8168 | -95.8012
2.8204 | -95.9025
2.8241 | -95.9859
2.8277 | -96.0414
2.8314 | -96.0655
2.8351 | -96.0385
2.8387 | -95.9707
2.8424 | -95.902
2.846 | -95.8416
2.8497 | -95.8132
2.8534 | -95.8313
2.857 | -95.8906
2.8607 | -95.9601
2.8643 | -96.0008
2.868 | -95.9924
2.8716 | -95.9421
2.8753 | -95.8811
2.879 | -95.8355
2.8826 | -95.8197
2.8863 | -95.8203
2.8899 | -95.8226
2.8936 | -95.8315
2.8973 | -95.8418
2.9009 | -95.8638
2.9046 | -95.9087
2.9082 | -95.9911
2.9119 | -96.0808
2.9155 | -96.1428
2.9192 | -96.1495
2.9229 | -96.1005
2.9265 | -96.0376
2.9302 | -96.0176
2.9338 | -96.0755
2.9375 | -96.2138
2.9411 | -96.4078
2.9448 | -96.6234
2.9485 | -96.8048
2.9521 | -96.9113
2.9558 | -96.9138
2.9594 | -96.839
2.9631 | -96.7148
2.9668 | -96.5763
2.9704 | -96.4727
2.9741 | -96.4356
2.9777 | -96.4705
2.9814 | -96.557
2.985 | -96.6559
2.9887 | -96.7467
2.9924 | -96.7951
2.996 | -96.7944
2.9997 | -96.7445
3.0033 | -96.6666
3.007 | -96.5654
3.0107 | -96.4579
3.0143 | -96.3807
3.018 | -96.3367
3.0216 | -96.3299
3.0253 | -96.3579
3.0289 | -96.4104
3.0326 | -96.4875
3.0363 | -96.5569
3.0399 | -96.6234
3.0436 | -96.6811
3.0472 | -96.7008
3.0509 | -96.6799
3.0546 | -96.6264
3.0582 | -96.5554
3.0619 | -96.4566
3.0655 | -96.3399
3.0692 | -96.219
3.0728 | -96.1243
3.0765 | -96.0992
3.0802 | -96.1665
3.0838 | -96.3298
3.0875 | -96.5506
3.0911 | -96.7747
3.0948 | -96.9378
3.0985 | -96.972
3.1021 | -96.8864
3.1058 | -96.7142
3.1094 | -96.5246
3.1131 | -96.3837
3.1167 | -96.3358
3.1204 | -96.3943
3.1241 | -96.5342
3.1277 | -96.7095
3.1314 | -96.8754
3.135 | -96.9901
3.1387 | -97.0541
3.1423 | -97.0937
3.146 | -97.1142
3.1497 | -97.0879
3.1533 | -97.0021
3.157 | -96.8748
3.1606 | -96.7361
3.1643 | -96.6446
3.168 | -96.5735
3.1716 | -96.5369
3.1753 | -96.5386
3.1789 | -96.5739
3.1826 | -96.6421
3.1862 | -96.744
3.1899 | -96.8715
3.1936 | -97.0002
3.1972 | -97.0999
3.2009 | -97.1434
3.2045 | -97.1379
3.2082 | -97.0933
3.2119 | -97.0246
3.2155 | -96.9683
3.2192 | -96.928
3.2228 | -96.9279
3.2265 | -96.9649
3.2301 | -97.0241
3.2338 | -97.0794
3.2375 | -97.1033
3.2411 | -97.0647
3.2448 | -96.94
3.2484 | -96.7579
3.2521 | -96.5582
3.2557 | -96.3847
3.2594 | -96.2537
3.2631 | -96.1906
3.2667 | -96.2186
3.2704 | -96.3122
3.274 | -96.4587
3.2777 | -96.6335
3.2814 | -96.8187
3.285 | -96.9773
3.2887 | -97.0829
3.2923 | -97.1243
3.296 | -97.1284
3.2996 | -97.1181
3.3033 | -97.1069
3.307 | -97.1114
3.3106 | -97.1349
3.3143 | -97.1747
3.3179 | -97.2162
3.3216 | -97.2432
3.3253 | -97.2649
3.3289 | -97.2699
3.3326 | -97.2496
3.3362 | -97.2216
3.3399 | -97.2071
3.3435 | -97.2176
3.3472 | -97.2567
3.3509 | -97.3257
3.3545 | -97.4223
3.3582 | -97.5126
3.3618 | -97.5866
3.3655 | -97.6362
3.3692 | -97.6654
3.3728 | -97.6844
3.3765 | -97.6971
3.3801 | -97.6904
3.3838 | -97.6717
3.3874 | -97.6321
3.3911 | -97.5518
3.3948 | -97.4363
3.3984 | -97.3064
3.4021 | -97.1744
3.4057 | -97.0645
3.4094 | -96.9917
3.4131 | -96.9581
3.4167 | -96.954
3.4204 | -96.9816
3.424 | -97.0292
3.4277 | -97.0951
3.4313 | -97.1783
3.435 | -97.266
3.4387 | -97.3273
3.4423 | -97.3383
3.446 | -97.3034
3.4496 | -97.2336
3.4533 | -97.1317
3.4569 | -97.017
3.4606 | -96.9001
3.4643 | -96.8043
3.4679 | -96.7596
3.4716 | -96.7726
3.4752 | -96.8425
3.4789 | -96.9596
3.4826 | -97.104
3.4862 | -97.2462
3.4899 | -97.3591
3.4935 | -97.4384
3.4972 | -97.4788
3.5008 | -97.4746
3.5045 | -97.4554
3.5082 | -97.4736
3.5118 | -97.5597
3.5155 | -97.7041
3.5191 | -97.8799
3.5228 | -98.047
3.5265 | -98.1635
3.5301 | -98.2065
3.5338 | -98.1785
3.5374 | -98.096
3.5411 | -98.0106
3.5447 | -97.9415
3.5484 | -97.8828
3.5521 | -97.8484
3.5557 | -97.8257
3.5594 | -97.8076
3.563 | -97.7821
3.5667 | -97.7516
3.5704 | -97.7213
3.574 | -97.6765
3.5777 | -97.6133
3.5813 | -97.5539
3.585 | -97.5191
3.5886 | -97.5172
3.5923 | -97.5558
3.596 | -97.6157
3.5996 | -97.7188
3.6033 | -97.8345
3.6069 | -97.9345
3.6106 | -98.0044
3.6142 | -98.0353
3.6179 | -98.0351
3.6216 | -98.0118
3.6252 | -97.997
3.6289 | -98.0181
3.6325 | -98.071
3.6362 | -98.143
3.6399 | -98.2097
3.6435 | -98.2362
3.6472 | -98.1954
3.6508 | -98.098
3.6545 | -97.9462
3.6581 | -97.7967
3.6618 | -97.673
3.6655 | -97.5846
3.6691 | -97.5374
3.6728 | -97.5314
3.6764 | -97.5562
3.6801 | -97.591
3.6838 | -97.6316
3.6874 | -97.6705
3.6911 | -97.6932
3.6947 | -97.7316
3.6984 | -97.7691
3.702 | -97.8295
3.7057 | -97.9116
3.7094 | -98.0128
3.713 | -98.122
3.7167 | -98.2174
3.7203 | -98.2923
3.724 | -98.3455
3.7277 | -98.3813
3.7313 | -98.4187
3.735 | -98.473
3.7386 | -98.5305
3.7423 | -98.5829
3.7459 | -98.6266
3.7496 | -98.6455
3.7533 | -98.6155
3.7569 | -98.5417
3.7606 | -98.4338
3.7642 | -98.3022
3.7679 | -98.1783
3.7715 | -98.0818
3.7752 | -98.0287
3.7789 | -98.0304
3.7825 | -98.076
3.7862 | -98.1621
3.7898 | -98.2797
3.7935 | -98.414
3.7972 | -98.5431
3.8008 | -98.6649
3.8045 | -98.7553
3.8081 | -98.7993
3.8118 | -98.8044
3.8154 | -98.7873
3.8191 | -98.75
3.8228 | -98.7079
3.8264 | -98.67
3.8301 | -98.6431
3.8337 | -98.6361
3.8374 | -98.6439
3.8411 | -98.6646
3.8447 | -98.6809
3.8484 | -98.6719
3.852 | -98.6471
3.8557 | -98.6111
3.8593 | -98.5575
3.863 | -98.483
3.8667 | -98.3995
3.8703 | -98.3139
3.874 | -98.237
3.8776 | -98.1807
3.8813 | -98.1633
3.885 | -98.1816
3.8886 | -98.2457
3.8923 | -98.3439
3.8959 | -98.4561
3.8996 | -98.564
3.9032 | -98.6454
3.9069 | -98.6911
3.9106 | -98.6976
3.9142 | -98.6705
3.9179 | -98.6406
3.9215 | -98.6341
3.9252 | -98.6576
3.9288 | -98.7074
3.9325 | -98.7834
3.9362 | -98.8766
3.9398 | -98.9747
3.9435 | -99.0599
3.9471 | -99.1292
3.9508 | -99.1635
3.9545 | -99.1472
3.9581 | -99.0803
3.9618 | -98.9674
3.9654 | -98.8258
3.9691 | -98.6831
3.9727 | -98.5622
3.9764 | -98.4668
3.9801 | -98.401
3.9837 | -98.3726
3.9874 | -98.3824
3.991 | -98.4265
3.9947 | -98.483
3.9984 | -98.5479
4.002 | -98.6152
4.0057 | -98.6767
4.0093 | -98.7205
4.013 | -98.7513
4.0166 | -98.7864
4.0203 | -98.829
4.024 | -98.8863
4.0276 | -98.96
4.0313 | -99.0558
4.0349 | -99.1516
4.0386 | -99.231
4.0423 | -99.2849
4.0459 | -99.2871
4.0496 | -99.2273
4.0532 | -99.1173
4.0569 | -98.9749
4.0605 | -98.8037
4.0642 | -98.6346
4.0679 | -98.4872
4.0715 | -98.3765
4.0752 | -98.3145
4.0788 | -98.31
4.0825 | -98.3611
4.0861 | -98.4634
4.0898 | -98.6061
4.0935 | -98.7626
4.0971 | -98.9148
4.1008 | -99.0514
4.1044 | -99.1464
4.1081 | -99.1994
4.1118 | -99.2142
4.1154 | -99.1984
4.1191 | -99.1827
4.1227 | -99.1723
4.1264 | -99.1625
4.13 | -99.1596
4.1337 | -99.1562
4.1374 | -99.155
4.141 | -99.1576
4.1447 | -99.1682
4.1483 | -99.1858
4.152 | -99.1923
4.1557 | -99.183
4.1593 | -99.1669
4.163 | -99.1264
4.1666 | -99.0691
4.1703 | -99.0124
4.1739 | -98.956
4.1776 | -98.9164
4.1813 | -98.8933
4.1849 | -98.8873
4.1886 | -98.8944
4.1922 | -98.9007
4.1959 | -98.9023
4.1996 | -98.8934
4.2032 | -98.8745
4.2069 | -98.8544
4.2105 | -98.8339
4.2142 | -98.8214
4.2178 | -98.8104
4.2215 | -98.7886
4.2252 | -98.7683
4.2288 | -98.7529
4.2325 | -98.7561
4.2361 | -98.7673
4.2398 | -98.7927
4.2435 | -98.8247
4.2471 | -98.8642
4.2508 | -98.9031
4.2544 | -98.9329
4.2581 | -98.9494
4.2617 | -98.9546
4.2654 | -98.9516
4.2691 | -98.95
4.2727 | -98.9519
4.2764 | -98.9745
4.28 | -99.0166
4.2837 | -99.0662
4.2873 | -99.1271
4.291 | -99.1927
4.2947 | -99.2517
4.2983 | -99.3035
4.302 | -99.3482
4.3056 | -99.3921
4.3093 | -99.4297
4.313 | -99.4652
4.3166 | -99.4978
4.3203 | -99.5342
4.3239 | -99.5789
4.3276 | -99.6258
4.3312 | -99.6631
4.3349 | -99.6923
4.3386 | -99.7162
4.3422 | -99.7354
4.3459 | -99.7433
4.3495 | -99.7435
4.3532 | -99.7416
4.3569 | -99.7373
4.3605 | -99.7283
4.3642 | -99.7171
4.3678 | -99.6985
4.3715 | -99.6707
4.3751 | -99.6378
4.3788 | -99.6063
4.3825 | -99.572
4.3861 | -99.5368
4.3898 | -99.508
4.3934 | -99.4763
4.3971 | -99.4564
4.4008 | -99.4407
4.4044 | -99.4249
4.4081 | -99.4079
4.4117 | -99.3944
4.4154 | -99.3894
4.419 | -99.4082
4.4227 | -99.4467
4.4264 | -99.4913
4.43 | -99.5519
4.4337 | -99.6188
4.4373 | -99.6888
4.441 | -99.7613
4.4446 | -99.8371
4.4483 | -99.9057
4.452 | -99.9647
4.4556 | -100.0037
4.4593 | -100.0228
4.4629 | -100.0207
4.4666 | -99.9925
4.4703 | -99.9395
4.4739 | -99.8738
4.4776 | -99.7995
4.4812 | -99.7262
4.4849 | -99.6632
4.4885 | -99.6177
4.4922 | -99.5962
4.4959 | -99.6057
4.4995 | -99.6357
4.5032 | -99.6762
4.5068 | -99.7202
4.5105 | -99.7521
4.5142 | -99.7655
4.5178 | -99.7669
4.5215 | -99.758
4.5251 | -99.7421
4.5288 | -99.7169
4.5324 | -99.7016
4.5361 | -99.6969
4.5398 | -99.7229
4.5434 | -99.763
4.5471 | -99.8268
4.5507 | -99.9055
4.5544 | -99.9864
4.558 | -100.0612
4.5617 | -100.1261
4.5654 | -100.167
4.569 | -100.1828
4.5727 | -100.1813
4.5763 | -100.1527
4.58 | -100.1031
4.5837 | -100.0451
4.5873 | -99.9822
4.591 | -99.9216
4.5946 | -99.8776
4.5983 | -99.8481
4.6019 | -99.8291
4.6056 | -99.8208
4.6093 | -99.8151
4.6129 | -99.8226
4.6166 | -99.8388
4.6202 | -99.8633
4.6239 | -99.8977
4.6276 | -99.9345
4.6312 | -99.9693
4.6349 | -99.9966
4.6385 | -100.0183
4.6422 | -100.0324
4.6458 | -100.036
4.6495 | -100.0327
4.6532 | -100.0169
4.6568 | -99.9925
4.6605 | -99.9639
4.6641 | -99.9251
4.6678 | -99.885
4.6715 | -99.8458
4.6751 | -99.8166
4.6788 | -99.7997
4.6824 | -99.7941
4.6861 | -99.7977
4.6897 | -99.8196
4.6934 | -99.8442
4.6971 | -99.8647
4.7007 | -99.8866
4.7044 | -99.902
4.708 | -99.9059
4.7117 | -99.903
4.7153 | -99.9079
4.719 | -99.9128
4.7227 | -99.9304
4.7263 | -99.958
4.73 | -99.9897
4.7336 | -100.0276
4.7373 | -100.0541
4.741 | -100.0669
4.7446 | -100.0589
4.7483 | -100.0256
4.7519 | -100.0107
4.7556 | -99.9989
4.7592 | -99.9823
4.7629 | -99.9623
4.7666 | -99.9506
4.7702 | -99.9472
4.7739 | -99.9409
4.7775 | -99.9441
4.7812 | -99.9589
4.7849 | -99.9796
4.7885 | -100.0063
4.7922 | -100.0478
4.7958 | -100.0963
4.7995 | -100.1486
4.8031 | -100.1993
4.8068 | -100.2356
4.8105 | -100.2718
4.8141 | -100.2964
4.8178 | -100.3101
4.8214 | -100.3145
4.8251 | -100.3087
4.8288 | -100.3018
4.8324 | -100.2986
4.8361 | -100.2998
4.8397 | -100.3031
4.8434 | -100.3024
4.847 | -100.2976
4.8507 | -100.2891
4.8544 | -100.2715
4.858 | -100.2538
4.8617 | -100.2374
4.8653 | -100.2145
4.869 | -100.1858
4.8727 | -100.1573
4.8763 | -100.1338
4.88 | -100.1207
4.8836 | -100.1207
4.8873 | -100.1198
4.8909 | -100.1338
4.8946 | -100.1598
4.8983 | -100.1849
4.9019 | -100.2122
4.9056 | -100.2382
4.9092 | -100.2551
4.9129 | -100.2658
4.9165 | -100.268
4.9202 | -100.2652
4.9239 | -100.2604
4.9275 | -100.2628
4.9312 | -100.2724
4.9348 | -100.2899
4.9385 | -100.3174
4.9422 | -100.3535
4.9458 | -100.3989
4.9495 | -100.4439
4.9531 | -100.4834
4.9568 | -100.5178
4.9604 | -100.5465
4.9641 | -100.5707
4.9678 | -100.5831
4.9714 | -100.585
4.9751 | -100.5863
4.9787 | -100.5842
4.9824 | -100.5626
4.9861 | -100.53
4.9897 | -100.4857
4.9934 | -100.4317
4.997 | -100.3712
5.0007 | -100.3093
5.0043 | -100.2529
5.008 | -100.2065
5.0117 | -100.1685
5.0153 | -100.1552
5.019 | -100.153
5.0226 | -100.1598
5.0263 | -100.185
5.03 | -100.2233
5.0336 | -100.2667
5.0373 | -100.3171
5.0409 | -100.3778
5.0446 | -100.4408
5.0482 | -100.5121
5.0519 | -100.585
5.0556 | -100.6605
5.0592 | -100.7421
5.0629 | -100.8167
5.0665 | -100.8863
5.0702 | -100.9435
5.0738 | -100.9964
5.0775 | -101.0429
5.0812 | -101.0763
5.0848 | -101.099
5.0885 | -101.106
5.0921 | -101.0943
5.0958 | -101.0807
5.0995 | -101.0543
5.1031 | -101.0151
5.1068 | -100.9808
5.1104 | -100.9481
5.1141 | -100.9053
5.1177 | -100.8524
5.1214 | -100.7919
5.1251 | -100.7145
5.1287 | -100.6315
5.1324 | -100.5489
5.136 | -100.4761
5.1397 | -100.4272
5.1434 | -100.4019
5.147 | -100.399
5.1507 | -100.4167
5.1543 | -100.4551
5.158 | -100.5007
5.1616 | -100.548
5.1653 | -100.5946
5.169 | -100.6368
5.1726 | -100.6694
5.1763 | -100.698
5.1799 | -100.7313
5.1836 | -100.7605
5.1873 | -100.7788
5.1909 | -100.7905
5.1946 | -100.7944
5.1982 | -100.7933
5.2019 | -100.7963
5.2055 | -100.7957
5.2092 | -100.8017
5.2129 | -100.8177
5.2165 | -100.8303
5.2202 | -100.8383
5.2238 | -100.8454
5.2275 | -100.842
5.2311 | -100.8286
5.2348 | -100.8277
5.2385 | -100.8327
5.2421 | -100.8488
5.2458 | -100.8776
5.2494 | -100.9255
5.2531 | -100.9858
5.2568 | -101.0439
5.2604 | -101.1043
5.2641 | -101.1756
5.2677 | -101.2411
5.2714 | -101.297
5.275 | -101.3043
5.2787 | -101.3213
5.2824 | -101.324
5.286 | -101.3164
5.2897 | -101.2878
5.2933 | -101.2471
5.297 | -101.2089
5.3007 | -101.1694
5.3043 | -101.1318
5.308 | -101.0931
5.3116 | -101.0544
5.3153 | -101.0203
5.3189 | -100.9947
5.3226 | -100.9723
5.3263 | -100.955
5.3299 | -100.9384
5.3336 | -100.9219
5.3372 | -100.9054
5.3409 | -100.8953
5.3446 | -100.8862
5.3482 | -100.8778
5.3519 | -100.8762
5.3555 | -100.882
5.3592 | -100.8959
5.3628 | -100.91
5.3665 | -100.9177
5.3702 | -100.9241
5.3738 | -100.9336
5.3775 | -100.9415
5.3811 | -100.9354
5.3848 | -100.9213
5.3884 | -100.902
5.3921 | -100.8774
5.3958 | -100.848
5.3994 | -100.8049
5.4031 | -100.7663
5.4067 | -100.7309
5.4104 | -100.7032
5.4141 | -100.6885
5.4177 | -100.6885
5.4214 | -100.7024
5.425 | -100.7264
5.4287 | -100.7657
5.4323 | -100.8126
5.436 | -100.8549
5.4397 | -100.8946
5.4433 | -100.9296
5.447 | -100.9671
5.4506 | -101.0012
5.4543 | -101.022
5.458 | -101.0297
5.4616 | -101.0385
5.4653 | -101.055
5.4689 | -101.0759
5.4726 | -101.0953
5.4762 | -101.1225
5.4799 | -101.1441
5.4836 | -101.1691
5.4872 | -101.1905
5.4909 | -101.2193
5.4945 | -101.2584
5.4982 | -101.3011
5.5019 | -101.3519
5.5055 | -101.4158
5.5092 | -101.4764
5.5128 | -101.5229
5.5165 | -101.5544
5.5201 | -101.5714
5.5238 | -101.5688
5.5275 | -101.5335
5.5311 | -101.4683
5.5348 | -101.3977
5.5384 | -101.332
5.5421 | -101.264
5.5457 | -101.186
5.5494 | -101.1171
5.5531 | -101.0598
5.5567 | -101.014
5.5604 | -100.9772
5.564 | -100.9591
5.5677 | -100.9504
5.5714 | -100.9592
5.575 | -100.9806
5.5787 | -101.0118
5.5823 | -101.0598
5.586 | -101.1168
5.5896 | -101.1739
5.5933 | -101.2381
5.597 | -101.307
5.6006 | -101.3695
5.6043 | -101.4175
5.6079 | -101.4626
5.6116 | -101.5098
5.6153 | -101.551
5.6189 | -101.5721
5.6226 | -101.5834
5.6262 | -101.5978
5.6299 | -101.6048
5.6335 | -101.591
5.6372 | -101.5647
5.6409 | -101.5226
5.6445 | -101.4732
5.6482 | -101.4137
5.6518 | -101.3458
5.6555 | -101.2717
5.6592 | -101.2029
5.6628 | -101.1417
5.6665 | -101.0954
5.6701 | -101.0641
5.6738 | -101.0514
5.6774 | -101.0539
5.6811 | -101.0692
5.6848 | -101.0984
5.6884 | -101.1341
5.6921 | -101.1713
5.6957 | -101.2171
5.6994 | -101.2704
5.7031 | -101.3279
5.7067 | -101.3795
5.7104 | -101.414
5.714 | -101.4329
5.7177 | -101.4318
5.7213 | -101.4024
5.725 | -101.3518
5.7287 | -101.2748
5.7323 | -101.1811
5.736 | -101.082
5.7396 | -100.985
5.7433 | -100.8938
5.7469 | -100.8259
5.7506 | -100.7733
5.7543 | -100.7467
5.7579 | -100.7558
5.7616 | -100.7937
5.7652 | -100.8559
5.7689 | -100.9419
5.7726 | -101.0444
5.7762 | -101.1517
5.7799 | -101.2561
5.7835 | -101.3532
5.7872 | -101.4348
5.7908 | -101.5091
5.7945 | -101.5711
5.7982 | -101.6232
5.8018 | -101.6419
5.8055 | -101.6719
5.8091 | -101.6893
5.8128 | -101.6979
5.8165 | -101.6946
5.8201 | -101.6816
5.8238 | -101.67
5.8274 | -101.6524
5.8311 | -101.6328
5.8347 | -101.6079
5.8384 | -101.5858
5.8421 | -101.5655
5.8457 | -101.5358
5.8494 | -101.5112
5.853 | -101.5011
5.8567 | -101.4943
5.8604 | -101.4949
5.864 | -101.5049
5.8677 | -101.5227
5.8713 | -101.5448
5.875 | -101.5715
5.8786 | -101.5837
5.8823 | -101.5873
5.886 | -101.5819
5.8896 | -101.5735
5.8933 | -101.5528
5.8969 | -101.5297
5.9006 | -101.5068
5.9042 | -101.4867
5.9079 | -101.4729
5.9116 | -101.47
5.9152 | -101.473
5.9189 | -101.489
5.9225 | -101.5164
5.9262 | -101.5511
5.9299 | -101.597
5.9335 | -101.6431
5.9372 | -101.6965
5.9408 | -101.7492
5.9445 | -101.7937
5.9481 | -101.8301
5.9518 | -101.8645
5.9555 | -101.8908
5.9591 | -101.9022
5.9628 | -101.9092
5.9664 | -101.9072
5.9701 | -101.9032
5.9738 | -101.8959
5.9774 | -101.8853
5.9811 | -101.8741
5.9847 | -101.8701
5.9884 | -101.8734
5.992 | -101.8785
5.9957 | -101.8881
5.9994 | -101.9033
6.003 | -101.9136
6.0067 | -101.9273
6.0103 | -101.9424
6.014 | -101.9521
6.0177 | -101.9531
6.0213 | -101.9481
6.025 | -101.9365
6.0286 | -101.9182
6.0323 | -101.8938
6.0359 | -101.8617
6.0396 | -101.8419
6.0433 | -101.8184
6.0469 | -101.7972
6.0506 | -101.7873
6.0542 | -101.7826
6.0579 | -101.7943
6.0615 | -101.8117
6.0652 | -101.8321
6.0689 | -101.8534
6.0725 | -101.87
6.0762 | -101.884
6.0798 | -101.8926
6.0835 | -101.8979
6.0872 | -101.9051
6.0908 | -101.9205
6.0945 | -101.9455
6.0981 | -101.9766
6.1018 | -102.0231
6.1054 | -102.0882
6.1091 | -102.1646
6.1128 | -102.2344
6.1164 | -102.3043
6.1201 | -102.3724
6.1237 | -102.4301
6.1274 | -102.4776
6.1311 | -102.4978
6.1347 | -102.5
6.1384 | -102.4833
6.142 | -102.4418
6.1457 | -102.3788
6.1493 | -102.3006
6.153 | -102.225
6.1567 | -102.1487
6.1603 | -102.0793
6.164 | -102.0254
6.1676 | -101.9831
6.1713 | -101.9499
6.175 | -101.9333
6.1786 | -101.9265
6.1823 | -101.9258
6.1859 | -101.935
6.1896 | -101.9503
6.1932 | -101.9642
6.1969 | -101.9813
6.2006 | -102.0125
6.2042 | -102.0515
6.2079 | -102.0879
6.2115 | -102.1171
6.2152 | -102.1426
6.2188 | -102.1572
6.2225 | -102.1652
6.2262 | -102.1617
6.2298 | -102.1425
6.2335 | -102.1069
6.2371 | -102.0662
6.2408 | -102.0182
6.2445 | -101.9642
6.2481 | -101.9109
6.2518 | -101.8581
6.2554 | -101.8032
6.2591 | -101.7519
6.2627 | -101.713
6.2664 | -101.6803
6.2701 | -101.6575
6.2737 | -101.6482
6.2774 | -101.6503
6.281 | -101.6648
6.2847 | -101.6976
6.2884 | -101.7379
6.292 | -101.7808
6.2957 | -101.8379
6.2993 | -101.8993
6.303 | -101.9635
6.3066 | -102.0288
6.3103 | -102.0892
6.314 | -102.1404
6.3176 | -102.1822
6.3213 | -102.2112
6.3249 | -102.2305
6.3286 | -102.2316
6.3323 | -102.2247
6.3359 | -102.2245
6.3396 | -102.2179
6.3432 | -102.2124
6.3469 | -102.2136
6.3505 | -102.2245
6.3542 | -102.2382
6.3579 | -102.2478
6.3615 | -102.2598
6.3652 | -102.2743
6.3688 | -102.2891
6.3725 | -102.2983
6.3761 | -102.3079
6.3798 | -102.3143
6.3835 | -102.3177
6.3871 | -102.3176
6.3908 | -102.3192
6.3944 | -102.3268
6.3981 | -102.3373
6.4018 | -102.3471
6.4054 | -102.3575
6.4091 | -102.3703
6.4127 | -102.3834
6.4164 | -102.3967
6.42 | -102.4009
6.4237 | -102.3976
6.4274 | -102.3794
6.431 | -102.3539
6.4347 | -102.3299
6.4383 | -102.3076
6.442 | -102.2758
6.4457 | -102.2393
6.4493 | -102.2013
6.453 | -102.168
6.4566 | -102.1383
6.4603 | -102.1094
6.4639 | -102.083
6.4676 | -102.0563
6.4713 | -102.036
6.4749 | -102.0203
6.4786 | -102.0165
6.4822 | -102.0227
6.4859 | -102.0344
6.4896 | -102.0452
6.4932 | -102.0659
6.4969 | -102.093
6.5005 | -102.1278
6.5042 | -102.1624
6.5078 | -102.1959
6.5115 | -102.2262
6.5152 | -102.2506
6.5188 | -102.2702
6.5225 | -102.2914
6.5261 | -102.3103
6.5298 | -102.3248
6.5334 | -102.341
6.5371 | -102.3553
6.5408 | -102.3754
6.5444 | -102.3929
6.5481 | -102.4132
6.5517 | -102.4337
6.5554 | -102.4545
6.5591 | -102.4778
6.5627 | -102.4997
6.5664 | -102.5276
6.57 | -102.557
6.5737 | -102.5859
6.5773 | -102.608
6.581 | -102.6357
6.5847 | -102.6568
6.5883 | -102.6722
6.592 | -102.674
6.5956 | -102.6612
6.5993 | -102.6362
6.603 | -102.6137
6.6066 | -102.5777
6.6103 | -102.536
6.6139 | -102.4962
6.6176 | -102.455
6.6212 | -102.4062
6.6249 | -102.364
6.6286 | -102.3216
6.6322 | -102.2803
6.6359 | -102.2489
6.6395 | -102.2258
6.6432 | -102.2118
6.6469 | -102.2057
6.6505 | -102.203
6.6542 | -102.2086
6.6578 | -102.2246
6.6615 | -102.2505
6.6651 | -102.2829
6.6688 | -102.3231
6.6725 | -102.3675
6.6761 | -102.4128
6.6798 | -102.4606
6.6834 | -102.5134
6.6871 | -102.5663
6.6907 | -102.6146
6.6944 | -102.6601
6.6981 | -102.6968
6.7017 | -102.7227
6.7054 | -102.7517
6.709 | -102.7711
6.7127 | -102.7885
6.7164 | -102.8074
6.72 | -102.8226
6.7237 | -102.8297
6.7273 | -102.8357
6.731 | -102.8445
6.7346 | -102.8464
6.7383 | -102.8448
6.742 | -102.8388
6.7456 | -102.8276
6.7493 | -102.8159
6.7529 | -102.7922
6.7566 | -102.7614
6.7603 | -102.7272
6.7639 | -102.6926
6.7676 | -102.6589
6.7712 | -102.6319
6.7749 | -102.6128
6.7785 | -102.5946
6.7822 | -102.5822
6.7859 | -102.5738
6.7895 | -102.5677
6.7932 | -102.567
6.7968 | -102.5668
6.8005 | -102.5685
6.8042 | -102.571
6.8078 | -102.5772
6.8115 | -102.5818
6.8151 | -102.5883
6.8188 | -102.6048
6.8224 | -102.6358
6.8261 | -102.6627
6.8298 | -102.6858
6.8334 | -102.7191
6.8371 | -102.745
6.8407 | -102.7661
6.8444 | -102.7756
6.848 | -102.7769
6.8517 | -102.7727
6.8554 | -102.7547
6.859 | -102.7166
6.8627 | -102.6989
6.8663 | -102.6797
6.87 | -102.6627
6.8737 | -102.6498
6.8773 | -102.6359
6.881 | -102.623
6.8846 | -102.608
6.8883 | -102.5923
6.8919 | -102.5776
6.8956 | -102.5587
6.8993 | -102.538
6.9029 | -102.5142
6.9066 | -102.4855
6.9102 | -102.4539
6.9139 | -102.4199
6.9176 | -102.381
6.9212 | -102.3397
6.9249 | -102.2937
6.9285 | -102.2559
6.9322 | -102.2193
6.9358 | -102.1823
6.9395 | -102.1597
6.9432 | -102.1477
6.9468 | -102.1508
6.9505 | -102.1655
6.9541 | -102.1889
6.9578 | -102.2209
6.9615 | -102.2563
6.9651 | -102.2956
6.9688 | -102.3289
6.9724 | -102.3588
6.9761 | -102.3868
6.9797 | -102.4125
6.9834 | -102.4405
6.9871 | -102.4638
6.9907 | -102.4858
6.9944 | -102.5014
6.998 | -102.5113
7.0017 | -102.519
7.0053 | -102.5244
7.009 | -102.5263
7.0127 | -102.5289
7.0163 | -102.5309
7.02 | -102.5252
7.0236 | -102.5214
7.0273 | -102.5161
7.031 | -102.5142
7.0346 | -102.5107
7.0383 | -102.5103
7.0419 | -102.5125
7.0456 | -102.5175
7.0492 | -102.5244
7.0529 | -102.5374
7.0566 | -102.5552
7.0602 | -102.5734
7.0639 | -102.5967
7.0675 | -102.622
7.0712 | -102.6492
7.0749 | -102.6791
7.0785 | -102.7044
7.0822 | -102.72
7.0858 | -102.7296
7.0895 | -102.7365
7.0931 | -102.7393
7.0968 | -102.7363
7.1005 | -102.7368
7.1041 | -102.7342
7.1078 | -102.7309
7.1114 | -102.7313
7.1151 | -102.7386
7.1188 | -102.7512
7.1224 | -102.7684
7.1261 | -102.7955
7.1297 | -102.8293
7.1334 | -102.8691
7.137 | -102.9052
7.1407 | -102.9407
7.1444 | -102.9674
7.148 | -102.9867
7.1517 | -102.9925
7.1553 | -102.989
7.159 | -102.9781
7.1627 | -102.9546
7.1663 | -102.9217
7.17 | -102.8841
7.1736 | -102.8457
7.1773 | -102.8076
7.1809 | -102.7719
7.1846 | -102.7432
7.1883 | -102.7222
7.1919 | -102.7043
7.1956 | -102.6974
7.1992 | -102.6968
7.2029 | -102.7048
7.2065 | -102.7164
7.2102 | -102.7369
7.2139 | -102.7675
7.2175 | -102.8005
7.2212 | -102.8323
7.2248 | -102.8719
7.2285 | -102.9144
7.2322 | -102.9519
7.2358 | -102.9805
7.2395 | -103.0064
7.2431 | -103.0328
7.2468 | -103.0589
7.2504 | -103.077
7.2541 | -103.0888
7.2578 | -103.0938
7.2614 | -103.0929
7.2651 | -103.083
7.2687 | -103.0727
7.2724 | -103.0593
7.2761 | -103.0337
7.2797 | -102.9989
7.2834 | -102.9557
7.287 | -102.9165
7.2907 | -102.8777
7.2943 | -102.8372
7.298 | -102.8004
7.3017 | -102.7648
7.3053 | -102.739
7.309 | -102.7175
7.3126 | -102.7031
7.3163 | -102.6985
7.32 | -102.6999
7.3236 | -102.716
7.3273 | -102.7431
7.3309 | -102.7729
7.3346 | -102.8122
7.3382 | -102.8592
7.3419 | -102.9131
7.3456 | -102.9625
7.3492 | -103.0077
7.3529 | -103.0469
7.3565 | -103.0851
7.3602 | -103.1164
7.3638 | -103.1411
7.3675 | -103.161
7.3712 | -103.1751
7.3748 | -103.1886
7.3785 | -103.198
7.3821 | -103.2039
7.3858 | -103.2166
7.3895 | -103.2273
7.3931 | -103.2342
7.3968 | -103.239
7.4004 | -103.2479
7.4041 | -103.2571
7.4077 | -103.2631
7.4114 | -103.2666
7.4151 | -103.2625
7.4187 | -103.2611
7.4224 | -103.2544
7.426 | -103.2462
7.4297 | -103.2324
7.4334 | -103.2188
7.437 | -103.2007
7.4407 | -103.1872
7.4443 | -103.1741
7.448 | -103.1504
7.4516 | -103.1311
7.4553 | -103.1151
7.459 | -103.0982
7.4626 | -103.0792
7.4663 | -103.0581
7.4699 | -103.0398
7.4736 | -103.0255
7.4773 | -103.0086
7.4809 | -102.9909
7.4846 | -102.977
7.4882 | -102.9729
7.4919 | -102.9699
7.4955 | -102.9715
7.4992 | -102.9733
7.5029 | -102.9847
7.5065 | -103.0033
7.5102 | -103.0185
7.5138 | -103.0364
7.5175 | -103.0574
7.5211 | -103.0807
7.5248 | -103.104
7.5285 | -103.1328
7.5321 | -103.1564
7.5358 | -103.1803
7.5394 | -103.2021
7.5431 | -103.225
7.5468 | -103.2437
7.5504 | -103.2662
7.5541 | -103.2856
7.5577 | -103.305
7.5614 | -103.3292
7.565 | -103.3333
7.5687 | -103.3358
7.5724 | -103.339
7.576 | -103.3359
7.5797 | -103.324
7.5833 | -103.3017
7.587 | -103.2744
7.5907 | -103.2429
7.5943 | -103.2053
7.598 | -103.1577
7.6016 | -103.1047
7.6053 | -103.056
7.6089 | -103.0063
7.6126 | -102.966
7.6163 | -102.9319
7.6199 | -102.9073
7.6236 | -102.8922
7.6272 | -102.8877
7.6309 | -102.8931
7.6346 | -102.9082
7.6382 | -102.9305
7.6419 | -102.9524
7.6455 | -102.9739
7.6492 | -102.9915
7.6528 | -103.0048
7.6565 | -103.0177
7.6602 | -103.021
7.6638 | -103.0116
7.6675 | -102.9984
7.6711 | -102.9803
7.6748 | -102.9591
7.6784 | -102.9399
7.6821 | -102.9141
7.6858 | -102.8877
7.6894 | -102.8711
7.6931 | -102.8622
7.6967 | -102.8558
7.7004 | -102.8565
7.7041 | -102.865
7.7077 | -102.8774
7.7114 | -102.8924
7.715 | -102.9091
7.7187 | -102.9343
7.7223 | -102.9608
7.726 | -102.9837
7.7297 | -103.0077
7.7333 | -103.0308
7.737 | -103.0595
7.7406 | -103.0834
7.7443 | -103.1047
7.748 | -103.1187
7.7516 | -103.1331
7.7553 | -103.1458
7.7589 | -103.1575
7.7626 | -103.1708
7.7662 | -103.1827
7.7699 | -103.1906
7.7736 | -103.2055
7.7772 | -103.2204
7.7809 | -103.2289
7.7845 | -103.2418
7.7882 | -103.256
7.7919 | -103.2681
7.7955 | -103.2786
7.7992 | -103.2876
7.8028 | -103.2925
7.8065 | -103.2998
7.8101 | -103.3123
7.8138 | -103.3203
7.8175 | -103.3305
7.8211 | -103.343
7.8248 | -103.3537
7.8284 | -103.3585
7.8321 | -103.3511
7.8357 | -103.3443
7.8394 | -103.339
7.8431 | -103.3315
7.8467 | -103.3266
7.8504 | -103.32
7.854 | -103.3203
7.8577 | -103.333
7.8614 | -103.3551
7.865 | -103.3787
7.8687 | -103.4104
7.8723 | -103.4409
7.876 | -103.4708
7.8796 | -103.5035
7.8833 | -103.5379
7.887 | -103.5663
7.8906 | -103.5915
7.8943 | -103.6081
7.8979 | -103.6203
7.9016 | -103.6317
7.9053 | -103.6357
7.9089 | -103.628
7.9126 | -103.6116
7.9162 | -103.5574
7.9199 | -103.5265
7.9235 | -103.4902
7.9272 | -103.4473
7.9309 | -103.403
7.9345 | -103.3548
7.9382 | -103.3082
7.9418 | -103.2629
7.9455 | -103.2238
7.9492 | -103.194
7.9528 | -103.171
7.9565 | -103.1569
7.9601 | -103.1496
7.9638 | -103.1536
7.9674 | -103.1679
7.9711 | -103.1891
7.9748 | -103.2136
7.9784 | -103.2461
7.9821 | -103.2824
7.9857 | -103.3161
7.9894 | -103.3518
7.9931 | -103.3882
7.9967 | -103.4238
8.0004 | -103.4549
8.004 | -103.4809
8.0077 | -103.5073
8.0113 | -103.5309
8.015 | -103.5403
8.0187 | -103.5449
8.0223 | -103.5516
8.026 | -103.5575
8.0296 | -103.555
8.0333 | -103.5526
8.0369 | -103.547
8.0406 | -103.5379
8.0443 | -103.529
8.0479 | -103.519
8.0516 | -103.5083
8.0552 | -103.5019
8.0589 | -103.4939
8.0626 | -103.4811
8.0662 | -103.4681
8.0699 | -103.4551
8.0735 | -103.4373
8.0772 | -103.4161
8.0808 | -103.3956
8.0845 | -103.3673
8.0882 | -103.3365
8.0918 | -103.3022
8.0955 | -103.2715
8.0991 | -103.2448
8.1028 | -103.2217
8.1065 | -103.2072
8.1101 | -103.1965
8.1138 | -103.191
8.1174 | -103.1929
8.1211 | -103.2013
8.1247 | -103.2161
8.1284 | -103.2344
8.1321 | -103.2589
8.1357 | -103.2879
8.1394 | -103.3115
8.143 | -103.3364
8.1467 | -103.3626
8.1503 | -103.392
8.154 | -103.4198
8.1577 | -103.444
8.1613 | -103.463
8.165 | -103.4767
8.1686 | -103.4919
8.1723 | -103.5055
8.176 | -103.51
8.1796 | -103.5121
8.1833 | -103.5192
8.1869 | -103.5229
8.1906 | -103.5243
8.1942 | -103.5299
8.1979 | -103.5331
8.2016 | -103.5385
8.2052 | -103.5501
8.2089 | -103.5632
8.2125 | -103.58
8.2162 | -103.5946
8.2199 | -103.6085
8.2235 | -103.6208
8.2272 | -103.6345
8.2308 | -103.6494
8.2345 | -103.6573
8.2381 | -103.6621
8.2418 | -103.6697
8.2455 | -103.6748
8.2491 | -103.6737
8.2528 | -103.6631
8.2564 | -103.6478
8.2601 | -103.6335
8.2638 | -103.6084
8.2674 | -103.5807
8.2711 | -103.5542
8.2747 | -103.5327
8.2784 | -103.5124
8.282 | -103.5005
8.2857 | -103.4904
8.2894 | -103.4799
8.293 | -103.4748
8.2967 | -103.4753
8.3003 | -103.4821
8.304 | -103.4829
8.3077 | -103.4819
8.3113 | -103.4816
8.315 | -103.487
8.3186 | -103.497
8.3223 | -103.5046
8.3259 | -103.511
8.3296 | -103.5228
8.3333 | -103.5344
8.3369 | -103.548
8.3406 | -103.5627
8.3442 | -103.5804
8.3479 | -103.5973
8.3515 | -103.6108
8.3552 | -103.6231
8.3589 | -103.6335
8.3625 | -103.6382
8.3662 | -103.6423
8.3698 | -103.6507
8.3735 | -103.6605
8.3772 | -103.6713
8.3808 | -103.6812
8.3845 | -103.6977
8.3881 | -103.7102
8.3918 | -103.7158
8.3954 | -103.7208
8.3991 | -103.7292
8.4028 | -103.7381
8.4064 | -103.7501
8.4101 | -103.7643
8.4137 | -103.7771
8.4174 | -103.7915
8.4211 | -103.8102
8.4247 | -103.8354
8.4284 | -103.857
8.432 | -103.8755
8.4357 | -103.8866
8.4393 | -103.8964
8.443 | -103.8973
8.4467 | -103.9008
8.4503 | -103.8981
8.454 | -103.8892
8.4576 | -103.8756
8.4613 | -103.8534
8.4649 | -103.8287
8.4686 | -103.8038
8.4723 | -103.7738
8.4759 | -103.7445
8.4796 | -103.717
8.4832 | -103.6892
8.4869 | -103.6632
8.4906 | -103.6414
8.4942 | -103.6211
8.4979 | -103.609
8.5015 | -103.6048
8.5052 | -103.608
8.5088 | -103.6168
8.5125 | -103.6363
8.5162 | -103.6665
8.5198 | -103.7064
8.5235 | -103.7519
8.5271 | -103.8022
8.5308 | -103.8514
8.5345 | -103.8995
8.5381 | -103.9439
8.5418 | -103.9795
8.5454 | -104.0042
8.5491 | -104.0241
8.5527 | -104.0392
8.5564 | -104.0432
8.5601 | -104.0379
8.5637 | -104.0278
8.5674 | -104.0107
8.571 | -103.9936
8.5747 | -103.9745
8.5784 | -103.9604
8.582 | -103.9489
8.5857 | -103.9421
8.5893 | -103.9396
8.593 | -103.9358
8.5966 | -103.9318
8.6003 | -103.9325
8.604 | -103.9364
8.6076 | -103.9421
8.6113 | -103.948
8.6149 | -103.9587
8.6186 | -103.9663
8.6223 | -103.9744
8.6259 | -103.9806
8.6296 | -103.987
8.6332 | -103.9952
8.6369 | -104.004
8.6405 | -104.0134
8.6442 | -104.021
8.6479 | -104.0231
8.6515 | -104.0279
8.6552 | -104.0286
8.6588 | -104.0309
8.6625 | -104.0325
8.6661 | -104.0354
8.6698 | -104.0405
8.6735 | -104.0413
8.6771 | -104.045
8.6808 | -104.0522
8.6844 | -104.0633
8.6881 | -104.0755
8.6918 | -104.0866
8.6954 | -104.1007
8.6991 | -104.1178
8.7027 | -104.1346
8.7064 | -104.1474
8.71 | -104.1543
8.7137 | -104.1536
8.7174 | -104.1483
8.721 | -104.1392
8.7247 | -104.129
8.7283 | -104.1141
8.732 | -104.0919
8.7357 | -104.0655
8.7393 | -104.0329
8.743 | -103.9923
8.7466 | -103.9528
8.7503 | -103.9146
8.7539 | -103.8784
8.7576 | -103.8412
8.7613 | -103.802
8.7649 | -103.7659
8.7686 | -103.7355
8.7722 | -103.7083
8.7759 | -103.6899
8.7796 | -103.6797
8.7832 | -103.682
8.7869 | -103.6888
8.7905 | -103.7025
8.7942 | -103.7229
8.7978 | -103.7453
8.8015 | -103.7755
8.8052 | -103.8077
8.8088 | -103.8384
8.8125 | -103.8679
8.8161 | -103.8992
8.8198 | -103.9299
8.8234 | -103.9608
8.8271 | -103.9893
8.8308 | -104.0184
8.8344 | -104.0467
8.8381 | -104.0768
8.8417 | -104.1087
8.8454 | -104.1392
8.8491 | -104.1684
8.8527 | -104.1918
8.8564 | -104.2124
8.86 | -104.2339
8.8637 | -104.2575
8.8673 | -104.2826
8.871 | -104.3069
8.8747 | -104.3246
8.8783 | -104.3382
8.882 | -104.3523
8.8856 | -104.3647
8.8893 | -104.3775
8.893 | -104.3885
8.8966 | -104.3945
8.9003 | -104.4015
8.9039 | -104.405
8.9076 | -104.4041
8.9112 | -104.3954
8.9149 | -104.3919
8.9186 | -104.3909
8.9222 | -104.3873
8.9259 | -104.3866
8.9295 | -104.3877
8.9332 | -104.388
8.9369 | -104.3932
8.9405 | -104.3981
8.9442 | -104.3973
8.9478 | -104.3956
8.9515 | -104.3955
8.9551 | -104.3988
8.9588 | -104.4005
8.9625 | -104.3973
8.9661 | -104.3885
8.9698 | -104.3647
8.9734 | -104.3457
8.9771 | -104.3228
8.9807 | -104.2924
8.9844 | -104.2618
8.9881 | -104.2304
8.9917 | -104.1968
8.9954 | -104.1597
8.999 | -104.1238
9.0027 | -104.09
9.0064 | -104.0598
9.01 | -104.0313
9.0137 | -104.0046
9.0173 | -103.9807
9.021 | -103.9566
9.0246 | -103.9386
9.0283 | -103.9296
9.032 | -103.9202
9.0356 | -103.913
9.0393 | -103.9114
9.0429 | -103.9159
9.0466 | -103.9156
9.0503 | -103.9149
9.0539 | -103.918
9.0576 | -103.9247
9.0612 | -103.9292
9.0649 | -103.934
9.0685 | -103.9407
9.0722 | -103.9484
9.0759 | -103.9549
9.0795 | -103.9571
9.0832 | -103.9539
9.0868 | -103.9454
9.0905 | -103.9391
9.0942 | -103.9314
9.0978 | -103.9216
9.1015 | -103.9106
9.1051 | -103.8984
9.1088 | -103.8858
9.1124 | -103.8771
9.1161 | -103.873
9.1198 | -103.8658
9.1234 | -103.8545
9.1271 | -103.8445
9.1307 | -103.8377
9.1344 | -103.8342
9.138 | -103.8319
9.1417 | -103.8321
9.1454 | -103.8323
9.149 | -103.834
9.1527 | -103.8389
9.1563 | -103.8486
9.16 | -103.8519
9.1637 | -103.8531
9.1673 | -103.8648
9.171 | -103.8768
9.1746 | -103.8879
9.1783 | -103.9033
9.1819 | -103.9237
9.1856 | -103.9421
9.1893 | -103.9635
9.1929 | -103.987
9.1966 | -104.01
9.2002 | -104.0365
9.2039 | -104.0626
9.2076 | -104.0837
9.2112 | -104.1046
9.2149 | -104.128
9.2185 | -104.147
9.2222 | -104.1596
9.2258 | -104.1676
9.2295 | -104.1708
9.2332 | -104.1767
9.2368 | -104.1792
9.2405 | -104.1823
9.2441 | -104.1821
9.2478 | -104.179
9.2515 | -104.1785
9.2551 | -104.1839
9.2588 | -104.1969
9.2624 | -104.2093
9.2661 | -104.2192
9.2697 | -104.2315
9.2734 | -104.2452
9.2771 | -104.2588
9.2807 | -104.2731
9.2844 | -104.2869
9.288 | -104.3053
9.2917 | -104.3208
9.2953 | -104.3388
9.299 | -104.3575
9.3027 | -104.367
9.3063 | -104.374
9.31 | -104.3801
9.3136 | -104.3817
9.3173 | -104.3781
9.321 | -104.3733
9.3246 | -104.362
9.3283 | -104.3482
9.3319 | -104.3349
9.3356 | -104.3182
9.3392 | -104.3064
9.3429 | -104.3008
9.3466 | -104.2982
9.3502 | -104.2961
9.3539 | -104.298
9.3575 | -104.3084
9.3612 | -104.317
9.3649 | -104.3254
9.3685 | -104.3293
9.3722 | -104.3336
9.3758 | -104.336
9.3795 | -104.3396
9.3831 | -104.3377
9.3868 | -104.3271
9.3905 | -104.3094
9.3941 | -104.289
9.3978 | -104.2697
9.4014 | -104.2483
9.4051 | -104.2223
9.4088 | -104.2001
9.4124 | -104.1859
9.4161 | -104.1729
9.4197 | -104.1631
9.4234 | -104.1536
9.427 | -104.1501
9.4307 | -104.1495
9.4344 | -104.1492
9.438 | -104.1548
9.4417 | -104.1658
9.4453 | -104.1789
9.449 | -104.19
9.4527 | -104.2004
9.4563 | -104.2102
9.46 | -104.2234
9.4636 | -104.2317
9.4673 | -104.2395
9.4709 | -104.2456
9.4746 | -104.2481
9.4783 | -104.2504
9.4819 | -104.2506
9.4856 | -104.2497
9.4892 | -104.2547
9.4929 | -104.2606
9.4965 | -104.2759
9.5002 | -104.2916
9.5039 | -104.3085
9.5075 | -104.3267
9.5112 | -104.3459
9.5148 | -104.3648
9.5185 | -104.3845
9.5222 | -104.4064
9.5258 | -104.4349
9.5295 | -104.4645
9.5331 | -104.4933
9.5368 | -104.5227
9.5404 | -104.5546
9.5441 | -104.585
9.5478 | -104.6126
9.5514 | -104.6326
9.5551 | -104.6501
9.5587 | -104.6657
9.5624 | -104.6756
9.5661 | -104.6866
9.5697 | -104.6977
9.5734 | -104.7029
9.577 | -104.696
9.5807 | -104.6891
9.5843 | -104.6729
9.588 | -104.6567
9.5917 | -104.6381
9.5953 | -104.615
9.599 | -104.5896
9.6026 | -104.5691
9.6063 | -104.5475
9.6099 | -104.5253
9.6136 | -104.5025
9.6173 | -104.4863
9.6209 | -104.472
9.6246 | -104.4624
9.6282 | -104.4533
9.6319 | -104.444
9.6356 | -104.4414
9.6392 | -104.4443
9.6429 | -104.4477
9.6465 | -104.4559
9.6502 | -104.4654
9.6538 | -104.4737
9.6575 | -104.4869
9.6612 | -104.4988
9.6648 | -104.5116
9.6685 | -104.5226
9.6721 | -104.5327
9.6758 | -104.5377
9.6795 | -104.5406
9.6831 | -104.5406
9.6868 | -104.5408
9.6904 | -104.5397
9.6941 | -104.5331
9.6977 | -104.5242
9.7014 | -104.5119
9.7051 | -104.497
9.7087 | -104.479
9.7124 | -104.4612
9.716 | -104.447
9.7197 | -104.4412
9.7234 | -104.4336
9.727 | -104.4361
9.7307 | -104.4417
9.7343 | -104.4487
9.738 | -104.4612
9.7416 | -104.4769
9.7453 | -104.4893
9.749 | -104.5068
9.7526 | -104.5305
9.7563 | -104.556
9.7599 | -104.5772
9.7636 | -104.6003
9.7673 | -104.6166
9.7709 | -104.6277
9.7746 | -104.6417
9.7782 | -104.6492
9.7819 | -104.6486
9.7855 | -104.6472
9.7892 | -104.6378
9.7929 | -104.6274
9.7965 | -104.6215
9.8002 | -104.6111
9.8038 | -104.599
9.8075 | -104.5875
9.8111 | -104.5745
9.8148 | -104.5634
9.8185 | -104.5577
9.8221 | -104.5524
9.8258 | -104.552
9.8294 | -104.5493
9.8331 | -104.5414
9.8368 | -104.5431
9.8404 | -104.5461
9.8441 | -104.5452
9.8477 | -104.5437
9.8514 | -104.5367
9.855 | -104.5289
9.8587 | -104.5178
9.8624 | -104.5001
9.866 | -104.4802
9.8697 | -104.4531
9.8733 | -104.4232
9.877 | -104.3972
9.8807 | -104.3658
9.8843 | -104.3307
9.888 | -104.2963
9.8916 | -104.2594
9.8953 | -104.2193
9.8989 | -104.178
9.9026 | -104.139
9.9063 | -104.1022
9.9099 | -104.0732
9.9136 | -104.0482
9.9172 | -104.0291
9.9209 | -104.0178
9.9246 | -104.0146
9.9282 | -104.0185
9.9319 | -104.0372
9.9355 | -104.0623
9.9392 | -104.0927
9.9428 | -104.1288
9.9465 | -104.1774
9.9502 | -104.2314
9.9538 | -104.2909
9.9575 | -104.3512
9.9611 | -104.4092
9.9648 | -104.4701
9.9684 | -104.5289
9.9721 | -104.5878
9.9758 | -104.645
9.9794 | -104.6898
9.9831 | -104.7324
9.9867 | -104.7703
9.9904 | -104.8016
9.9941 | -104.8292
9.9977 | -104.8495
10.0014 | -104.8655
10.005 | -104.8706
10.0087 | -104.8666
10.0123 | -104.8639
10.016 | -104.8504
10.0197 | -104.8249
10.0233 | -104.8012
10.027 | -104.7751
10.0306 | -104.7548
10.0343 | -104.7357
10.038 | -104.7125
10.0416 | -104.6913
10.0453 | -104.6714
10.0489 | -104.651
10.0526 | -104.636
10.0562 | -104.6228
10.0599 | -104.6152
10.0636 | -104.6069
10.0672 | -104.6013
10.0709 | -104.5996
10.0745 | -104.6014
10.0782 | -104.6019
10.0819 | -104.6046
10.0855 | -104.6055
10.0892 | -104.6098
10.0928 | -104.6088
10.0965 | -104.6087
10.1001 | -104.6119
10.1038 | -104.6146
10.1075 | -104.6166
10.1111 | -104.6093
10.1148 | -104.5969
10.1184 | -104.5878
10.1221 | -104.5745
10.1257 | -104.5608
10.1294 | -104.5472
10.1331 | -104.5308
10.1367 | -104.5122
10.1404 | -104.494
10.144 | -104.4777
10.1477 | -104.4602
10.1514 | -104.4473
10.155 | -104.434
10.1587 | -104.4198
10.1623 | -104.4127
10.166 | -104.4114
10.1696 | -104.4094
10.1733 | -104.4084
10.177 | -104.4115
10.1806 | -104.4158
10.1843 | -104.4247
10.1879 | -104.4398
10.1916 | -104.4543
10.1953 | -104.4736
10.1989 | -104.4986
10.2026 | -104.522
10.2062 | -104.5485
10.2099 | -104.5783
10.2135 | -104.6106
10.2172 | -104.645
10.2209 | -104.6793
10.2245 | -104.7155
10.2282 | -104.751
10.2318 | -104.7844
10.2355 | -104.821
10.2392 | -104.8566
10.2428 | -104.8894
10.2465 | -104.9171
10.2501 | -104.9408
10.2538 | -104.962
10.2574 | -104.9787
10.2611 | -104.9961
10.2648 | -105.0116
10.2684 | -105.019
10.2721 | -105.0276
10.2757 | -105.0299
10.2794 | -105.0252
10.283 | -105.0223
10.2867 | -105.0188
10.2904 | -105.0098
10.294 | -104.9974
10.2977 | -104.9843
10.3013 | -104.9651
10.305 | -104.9436
10.3087 | -104.9209
10.3123 | -104.893
10.316 | -104.8603
10.3196 | -104.8221
10.3233 | -104.782
10.3269 | -104.7419
10.3306 | -104.7021
10.3343 | -104.663
10.3379 | -104.6238
10.3416 | -104.5868
10.3452 | -104.551
10.3489 | -104.5148
10.3526 | -104.485
10.3562 | -104.4598
10.3599 | -104.4354
10.3635 | -104.4148
10.3672 | -104.4001
10.3708 | -104.3941
10.3745 | -104.3955
10.3782 | -104.4038
10.3818 | -104.4179
10.3855 | -104.4364
10.3891 | -104.4578
10.3928 | -104.483
10.3965 | -104.5123
10.4001 | -104.5439
10.4038 | -104.5817
10.4074 | -104.6213
10.4111 | -104.6597
10.4147 | -104.6981
10.4184 | -104.7338
10.4221 | -104.7632
10.4257 | -104.7931
10.4294 | -104.8185
10.433 | -104.8375
10.4367 | -104.8544
10.4403 | -104.8686
10.444 | -104.8772
10.4477 | -104.882
10.4513 | -104.8852
10.455 | -104.8825
10.4586 | -104.8772
10.4623 | -104.8755
10.466 | -104.8738
10.4696 | -104.8669
10.4733 | -104.8589
10.4769 | -104.8481
10.4806 | -104.8396
10.4842 | -104.8267
10.4879 | -104.8172
10.4916 | -104.8078
10.4952 | -104.7967
10.4989 | -104.786
10.5025 | -104.7732
10.5062 | -104.7608
10.5099 | -104.7537
10.5135 | -104.7457
10.5172 | -104.739
10.5208 | -104.7299
10.5245 | -104.719
10.5281 | -104.7147
10.5318 | -104.7137
10.5355 | -104.7158
10.5391 | -104.7224
10.5428 | -104.7324
10.5464 | -104.7441
10.5501 | -104.7567
10.5538 | -104.7808
10.5574 | -104.7952
10.5611 | -104.81
10.5647 | -104.8256
10.5684 | -104.843
10.572 | -104.864
10.5757 | -104.886
10.5794 | -104.9005
10.583 | -104.9175
10.5867 | -104.9321
10.5903 | -104.9442
10.594 | -104.9574
10.5976 | -104.9694
10.6013 | -104.9817
10.605 | -104.9905
10.6086 | -104.9968
10.6123 | -104.9983
10.6159 | -104.994
10.6196 | -104.9856
10.6233 | -104.9848
10.6269 | -104.9813
10.6306 | -104.9753
10.6342 | -104.9615
10.6379 | -104.9493
10.6415 | -104.933
10.6452 | -104.9178
10.6489 | -104.9052
10.6525 | -104.8955
10.6562 | -104.8862
10.6598 | -104.8725
10.6635 | -104.8577
10.6672 | -104.8473
10.6708 | -104.8379
10.6745 | -104.8316
10.6781 | -104.8279
10.6818 | -104.8228
10.6854 | -104.8201
10.6891 | -104.8175
10.6928 | -104.8195
10.6964 | -104.8258
10.7001 | -104.8325
10.7037 | -104.8397
10.7074 | -104.849
10.7111 | -104.8625
10.7147 | -104.8807
10.7184 | -104.8991
10.722 | -104.9176
10.7257 | -104.9404
10.7293 | -104.9657
10.733 | -104.9904
10.7367 | -105.0136
10.7403 | -105.0363
10.744 | -105.0632
10.7476 | -105.0818
10.7513 | -105.092
10.7549 | -105.0925
10.7586 | -105.0934
10.7623 | -105.0973
10.7659 | -105.0927
10.7696 | -105.0839
10.7732 | -105.0709
10.7769 | -105.053
10.7806 | -105.0263
10.7842 | -104.993
10.7879 | -104.9561
10.7915 | -104.923
10.7952 | -104.8866
10.7988 | -104.8463
10.8025 | -104.8056
10.8062 | -104.7622
10.8098 | -104.718
10.8135 | -104.6744
10.8171 | -104.6417
10.8208 | -104.6158
10.8245 | -104.5965
10.8281 | -104.581
10.8318 | -104.5702
10.8354 | -104.5668
10.8391 | -104.5696
10.8427 | -104.5796
10.8464 | -104.5946
10.8501 | -104.612
10.8537 | -104.6363
10.8574 | -104.6674
10.861 | -104.6982
10.8647 | -104.7265
10.8684 | -104.7565
10.872 | -104.7859
10.8757 | -104.8151
10.8793 | -104.843
10.883 | -104.8652
10.8866 | -104.8804
10.8903 | -104.8936
10.894 | -104.9073
10.8976 | -104.9121
10.9013 | -104.9092
10.9049 | -104.9085
10.9086 | -104.908
10.9123 | -104.9116
10.9159 | -104.9053
10.9196 | -104.8958
10.9232 | -104.8906
10.9269 | -104.8834
10.9305 | -104.8775
10.9342 | -104.8759
10.9379 | -104.8712
10.9415 | -104.8697
10.9452 | -104.8674
10.9488 | -104.8656
10.9525 | -104.8663
10.9561 | -104.8681
10.9598 | -104.8703
10.9635 | -104.8769
10.9671 | -104.8842
10.9708 | -104.8918
10.9744 | -104.8955
10.9781 | -104.8985
10.9818 | -104.9002
10.9854 | -104.9046
10.9891 | -104.9142
10.9927 | -104.9239
10.9964 | -104.9333
11 | -104.9408
11.0037 | -104.9509
11.0074 | -104.965
11.011 | -104.9779
11.0147 | -104.9919
11.0183 | -105.0096
11.022 | -105.0269
11.0257 | -105.0397
11.0293 | -105.0502
11.033 | -105.0634
11.0366 | -105.072
11.0403 | -105.079
11.0439 | -105.088
11.0476 | -105.0988
11.0513 | -105.1063
11.0549 | -105.1106
11.0586 | -105.1161
11.0622 | -105.1213
11.0659 | -105.1221
11.0696 | -105.1175
11.0732 | -105.1143
11.0769 | -105.1071
11.0805 | -105.0974
11.0842 | -105.0924
11.0878 | -105.0865
11.0915 | -105.0742
11.0952 | -105.065
11.0988 | -105.0587
11.1025 | -105.0522
11.1061 | -105.0473
11.1098 | -105.0421
11.1134 | -105.0318
11.1171 | -105.0215
11.1208 | -105.0089
11.1244 | -104.9946
11.1281 | -104.984
11.1317 | -104.9744
11.1354 | -104.9641
11.1391 | -104.9528
11.1427 | -104.9445
11.1464 | -104.9367
11.15 | -104.9247
11.1537 | -104.9113
11.1573 | -104.9002
11.161 | -104.8895
11.1647 | -104.8823
11.1683 | -104.8719
11.172 | -104.8596
11.1756 | -104.848
11.1793 | -104.8401
11.183 | -104.834
11.1866 | -104.8299
11.1903 | -104.8266
11.1939 | -104.8241
11.1976 | -104.8251
11.2012 | -104.8257
11.2049 | -104.8268
11.2086 | -104.8292
11.2122 | -104.8326
11.2159 | -104.8364
11.2195 | -104.8402
11.2232 | -104.8469
11.2269 | -104.8582
11.2305 | -104.8702
11.2342 | -104.8805
11.2378 | -104.8919
11.2415 | -104.9057
11.2451 | -104.9199
11.2488 | -104.9338
11.2525 | -104.9517
11.2561 | -104.9719
11.2598 | -104.9877
11.2634 | -105.0063
11.2671 | -105.0243
11.2707 | -105.0415
11.2744 | -105.0581
11.2781 | -105.078
11.2817 | -105.0995
11.2854 | -105.1184
11.289 | -105.1335
11.2927 | -105.1458
11.2964 | -105.157
11.3 | -105.1636
11.3037 | -105.1663
11.3073 | -105.1632
11.311 | -105.1587
11.3146 | -105.1545
11.3183 | -105.1475
11.322 | -105.137
11.3256 | -105.1236
11.3293 | -105.1041
11.3329 | -105.0861
11.3366 | -105.0663
11.3403 | -105.0452
11.3439 | -105.0244
11.3476 | -105.0035
11.3512 | -104.9854
11.3549 | -104.9675
11.3585 | -104.9502
11.3622 | -104.9323
11.3659 | -104.9152
11.3695 | -104.9005
11.3732 | -104.8892
11.3768 | -104.8842
11.3805 | -104.8811
11.3842 | -104.8784
11.3878 | -104.8765
11.3915 | -104.8785
11.3951 | -104.8856
11.3988 | -104.895
11.4024 | -104.9064
11.4061 | -104.9152
11.4098 | -104.9259
11.4134 | -104.9382
11.4171 | -104.9507
11.4207 | -104.965
11.4244 | -104.9781
11.428 | -104.9851
11.4317 | -104.99
11.4354 | -104.9964
11.439 | -105.0007
11.4427 | -105.0052
11.4463 | -105.0041
11.45 | -105.0062
11.4537 | -105.0078
11.4573 | -105.004
11.461 | -104.9966
11.4646 | -104.9907
11.4683 | -104.9853
11.4719 | -104.9762
11.4756 | -104.9639
11.4793 | -104.9526
11.4829 | -104.941
11.4866 | -104.9241
11.4902 | -104.9047
11.4939 | -104.8874
11.4976 | -104.8701
11.5012 | -104.8542
11.5049 | -104.8401
11.5085 | -104.8273
11.5122 | -104.8179
11.5158 | -104.8061
11.5195 | -104.7948
11.5232 | -104.7863
11.5268 | -104.7799
11.5305 | -104.7768
11.5341 | -104.7722
11.5378 | -104.769
11.5415 | -104.7644
11.5451 | -104.7607
11.5488 | -104.7584
11.5524 | -104.7562
11.5561 | -104.7519
11.5597 | -104.7508
11.5634 | -104.7508
11.5671 | -104.7495
11.5707 | -104.7495
11.5744 | -104.7476
11.578 | -104.7474
11.5817 | -104.7483
11.5853 | -104.7504
11.589 | -104.7566
11.5927 | -104.7658
11.5963 | -104.7768
11.6 | -104.7894
11.6036 | -104.8051
11.6073 | -104.8322
11.611 | -104.8503
11.6146 | -104.8717
11.6183 | -104.8928
11.6219 | -104.9166
11.6256 | -104.9435
11.6292 | -104.9725
11.6329 | -105.0024
11.6366 | -105.0322
11.6402 | -105.0638
11.6439 | -105.0953
11.6475 | -105.1251
11.6512 | -105.1565
11.6549 | -105.192
11.6585 | -105.2247
11.6622 | -105.2509
11.6658 | -105.2719
11.6695 | -105.2933
11.6731 | -105.3111
11.6768 | -105.3231
11.6805 | -105.3296
11.6841 | -105.3346
11.6878 | -105.3342
11.6914 | -105.3326
11.6951 | -105.329
11.6988 | -105.3194
11.7024 | -105.3067
11.7061 | -105.2884
11.7097 | -105.2694
11.7134 | -105.2481
11.717 | -105.219
11.7207 | -105.1923
11.7244 | -105.1649
11.728 | -105.1317
11.7317 | -105.0978
11.7353 | -105.0615
11.739 | -105.0284
11.7426 | -104.9943
11.7463 | -104.9594
11.75 | -104.9217
11.7536 | -104.8851
11.7573 | -104.8473
11.7609 | -104.8134
11.7646 | -104.7798
11.7683 | -104.7457
11.7719 | -104.7129
11.7756 | -104.6826
11.7792 | -104.6546
11.7829 | -104.6272
11.7865 | -104.6039
11.7902 | -104.5829
11.7939 | -104.5638
11.7975 | -104.5503
11.8012 | -104.5406
11.8048 | -104.5324
11.8085 | -104.5282
11.8122 | -104.5265
11.8158 | -104.5285
11.8195 | -104.5333
11.8231 | -104.5416
11.8268 | -104.5502
11.8304 | -104.5627
11.8341 | -104.5778
11.8378 | -104.5981
11.8414 | -104.6195
11.8451 | -104.6416
11.8487 | -104.6647
11.8524 | -104.6895
11.8561 | -104.7191
11.8597 | -104.7509
11.8634 | -104.7761
11.867 | -104.7992
11.8707 | -104.8235
11.8743 | -104.8479
11.878 | -104.8697
11.8817 | -104.8884
11.8853 | -104.9087
11.889 | -104.9276
11.8926 | -104.9382
11.8963 | -104.9504
11.9 | -104.9629
11.9036 | -104.9716
11.9073 | -104.9772
11.9109 | -104.9779
11.9146 | -104.9806
11.9182 | -104.9794
11.9219 | -104.9782
11.9256 | -104.9764
11.9292 | -104.975
11.9329 | -104.9712
11.9365 | -104.9693
11.9402 | -104.9662
11.9438 | -104.9655
11.9475 | -104.9675
11.9512 | -104.9723
11.9548 | -104.9754
11.9585 | -104.9779
11.9621 | -104.9792
11.9658 | -104.9811
11.9695 | -104.9825
11.9731 | -104.9864
11.9768 | -104.9916
11.9804 | -104.995
11.9841 | -105.0006
11.9877 | -105.0038
11.9914 | -105.0048
11.9951 | -105.0056
11.9987 | -105.0087
12.0024 | -105.0132
12.006 | -105.0167
12.0097 | -105.0173
12.0134 | -105.0167
12.017 | -105.0179
12.0207 | -105.0198
12.0243 | -105.0197
12.028 | -105.0152
12.0316 | -105.0114
12.0353 | -105.0106
12.039 | -105.0093
12.0426 | -105.0095
12.0463 | -105.01
12.0499 | -105.0086
12.0536 | -105.0084
12.0572 | -105.0074
12.0609 | -105.0064
12.0646 | -105.0035
12.0682 | -104.9986
12.0719 | -104.9944
12.0755 | -104.9912
12.0792 | -104.9854
12.0829 | -104.9606
12.0865 | -104.9351
12.0902 | -104.9161
12.0938 | -104.8949
12.0975 | -104.8704
12.1011 | -104.8397
12.1048 | -104.8064
12.1085 | -104.7709
12.1121 | -104.7328
12.1158 | -104.6883
12.1194 | -104.6399
12.1231 | -104.5936
12.1268 | -104.5492
12.1304 | -104.5047
12.1341 | -104.4622
12.1377 | -104.4316
12.1414 | -104.3981
12.145 | -104.3675
12.1487 | -104.3388
12.1524 | -104.3167
12.156 | -104.3018
12.1597 | -104.2908
12.1633 | -104.2875
12.167 | -104.2906
12.1707 | -104.2999
12.1743 | -104.3152
12.178 | -104.3369
12.1816 | -104.3642
12.1853 | -104.3962
12.1889 | -104.4367
12.1926 | -104.4811
12.1963 | -104.5271
12.1999 | -104.5775
12.2036 | -104.6301
12.2072 | -104.682
12.2109 | -104.7335
12.2145 | -104.7832
12.2182 | -104.8283
12.2219 | -104.8717
12.2255 | -104.9134
12.2292 | -104.9508
12.2328 | -104.9846
12.2365 | -105.0179
12.2402 | -105.0463
12.2438 | -105.0679
12.2475 | -105.0836
12.2511 | -105.0962
12.2548 | -105.1113
12.2584 | -105.1352
12.2621 | -105.1466
12.2658 | -105.1464
12.2694 | -105.145
12.2731 | -105.1436
12.2767 | -105.1449
12.2804 | -105.1432
12.2841 | -105.1397
12.2877 | -105.138
12.2914 | -105.1395
12.295 | -105.1408
12.2987 | -105.1375
12.3023 | -105.1362
12.306 | -105.1364
12.3097 | -105.1363
12.3133 | -105.1342
12.317 | -105.1352
12.3206 | -105.138
12.3243 | -105.1403
12.328 | -105.1455
12.3316 | -105.1498
12.3353 | -105.1524
12.3389 | -105.1576
12.3426 | -105.1624
12.3462 | -105.1698
12.3499 | -105.1773
12.3536 | -105.1847
12.3572 | -105.1905
12.3609 | -105.2015
12.3645 | -105.2122
12.3682 | -105.2185
12.3719 | -105.2272
12.3755 | -105.2392
12.3792 | -105.2529
12.3828 | -105.2657
12.3865 | -105.2804
12.3901 | -105.2931
12.3938 | -105.3041
12.3975 | -105.314
12.4011 | -105.3211
12.4048 | -105.3275
12.4084 | -105.3325
12.4121 | -105.3356
12.4157 | -105.3354
12.4194 | -105.3323
12.4231 | -105.329
12.4267 | -105.3252
12.4304 | -105.3184
12.434 | -105.3083
12.4377 | -105.2988
12.4414 | -105.2882
12.445 | -105.2747
12.4487 | -105.2626
12.4523 | -105.2502
12.456 | -105.2334
12.4596 | -105.221
12.4633 | -105.2123
12.467 | -105.2045
12.4706 | -105.1946
12.4743 | -105.1832
12.4779 | -105.1735
12.4816 | -105.161
12.4853 | -105.1464
12.4889 | -105.1339
12.4926 | -105.1236
12.4962 | -105.1098
12.4999 | -105.094
12.5035 | -105.0834
12.5072 | -105.0723
12.5109 | -105.06
12.5145 | -105.0487
12.5182 | -105.0369
12.5218 | -105.0279
12.5255 | -105.019
12.5292 | -105.0127
12.5328 | -105.0048
12.5365 | -104.9981
12.5401 | -104.9902
12.5438 | -104.9861
12.5474 | -104.9848
12.5511 | -104.9812
12.5548 | -104.9771
12.5584 | -104.9724
12.5621 | -104.9681
12.5657 | -104.9649
12.5694 | -104.9635
12.573 | -104.9634
12.5767 | -104.9629
12.5804 | -104.9611
12.584 | -104.9582
12.5877 | -104.9559
12.5913 | -104.9505
12.595 | -104.9429
12.5987 | -104.9387
12.6023 | -104.9311
12.606 | -104.9214
12.6096 | -104.91
12.6133 | -104.8978
12.6169 | -104.8864
12.6206 | -104.8773
12.6243 | -104.8686
12.6279 | -104.8558
12.6316 | -104.8437
12.6352 | -104.8341
12.6389 | -104.8252
12.6426 | -104.8146
12.6462 | -104.8064
12.6499 | -104.8008
12.6535 | -104.7983
12.6572 | -104.8012
12.6608 | -104.8074
12.6645 | -104.8242
12.6682 | -104.8358
12.6718 | -104.8466
12.6755 | -104.8623
12.6791 | -104.879
12.6828 | -104.9005
12.6865 | -104.9225
12.6901 | -104.9481
12.6938 | -104.9773
12.6974 | -105.0062
12.7011 | -105.0378
12.7047 | -105.0664
12.7084 | -105.0969
12.7121 | -105.1298
12.7157 | -105.1655
12.7194 | -105.1989
12.723 | -105.2314
12.7267 | -105.2631
12.7303 | -105.2925
12.734 | -105.3218
12.7377 | -105.3496
12.7413 | -105.3735
12.745 | -105.3956
12.7486 | -105.4136
12.7523 | -105.4277
12.756 | -105.4407
12.7596 | -105.4508
12.7633 | -105.4569
12.7669 | -105.4619
12.7706 | -105.4602
12.7742 | -105.456
12.7779 | -105.4512
12.7816 | -105.4441
12.7852 | -105.4349
12.7889 | -105.4264
12.7925 | -105.416
12.7962 | -105.4021
12.7999 | -105.3838
12.8035 | -105.3621
12.8072 | -105.3435
12.8108 | -105.3269
12.8145 | -105.3086
12.8181 | -105.2895
12.8218 | -105.2694
12.8255 | -105.2515
12.8291 | -105.2352
12.8328 | -105.2201
12.8364 | -105.2075
12.8401 | -105.1942
12.8438 | -105.1845
12.8474 | -105.1787
12.8511 | -105.1737
12.8547 | -105.1706
12.8584 | -105.1722
12.862 | -105.1701
12.8657 | -105.1713
12.8694 | -105.1713
12.873 | -105.1701
12.8767 | -105.1737
12.8803 | -105.1777
12.884 | -105.1836
12.8876 | -105.1863
12.8913 | -105.1885
12.895 | -105.1874
12.8986 | -105.1839
12.9023 | -105.1805
12.9059 | -105.1762
12.9096 | -105.1693
12.9133 | -105.1596
12.9169 | -105.1477
12.9206 | -105.1351
12.9242 | -105.1224
12.9279 | -105.112
12.9315 | -105.1006
12.9352 | -105.0877
12.9389 | -105.0786
12.9425 | -105.0699
12.9462 | -105.061
12.9498 | -105.0545
12.9535 | -105.0489
12.9572 | -105.0448
12.9608 | -105.0424
12.9645 | -105.0419
12.9681 | -105.0417
12.9718 | -105.0454
12.9754 | -105.0475
12.9791 | -105.0517
12.9828 | -105.0605
12.9864 | -105.0701
12.9901 | -105.0813
12.9937 | -105.0938
12.9974 | -105.107
13.0011 | -105.1217
13.0047 | -105.1405
13.0084 | -105.1596
13.012 | -105.176
13.0157 | -105.1929
13.0193 | -105.2117
13.023 | -105.2298
13.0267 | -105.2505
13.0303 | -105.2721
13.034 | -105.2941
13.0376 | -105.3153
13.0413 | -105.3364
13.0449 | -105.3554
13.0486 | -105.3724
13.0523 | -105.3875
13.0559 | -105.4008
13.0596 | -105.4104
13.0632 | -105.4139
13.0669 | -105.4189
13.0706 | -105.424
13.0742 | -105.4288
13.0779 | -105.4311
13.0815 | -105.4286
13.0852 | -105.4252
13.0888 | -105.4213
13.0925 | -105.4184
13.0962 | -105.414
13.0998 | -105.4087
13.1035 | -105.4053
13.1071 | -105.3986
13.1108 | -105.3906
13.1145 | -105.3812
13.1181 | -105.3752
13.1218 | -105.3684
13.1254 | -105.3598
13.1291 | -105.3499
13.1327 | -105.3396
13.1364 | -105.333
13.1401 | -105.3229
13.1437 | -105.3122
13.1474 | -105.3009
13.151 | -105.2886
13.1547 | -105.2756
13.1584 | -105.2612
13.162 | -105.2471
13.1657 | -105.2336
13.1693 | -105.2183
13.173 | -105.208
13.1766 | -105.1976
13.1803 | -105.1834
13.184 | -105.1665
13.1876 | -105.1523
13.1913 | -105.1508
13.1949 | -105.1453
13.1986 | -105.1437
13.2022 | -105.1447
13.2059 | -105.143
13.2096 | -105.1433
13.2132 | -105.1444
13.2169 | -105.144
13.2205 | -105.1414
13.2242 | -105.1415
13.2279 | -105.1407
13.2315 | -105.1402
13.2352 | -105.142
13.2388 | -105.1465
13.2425 | -105.1512
13.2461 | -105.1565
13.2498 | -105.1617
13.2535 | -105.1638
13.2571 | -105.1634
13.2608 | -105.1639
13.2644 | -105.1646
13.2681 | -105.1677
13.2718 | -105.1746
13.2754 | -105.1795
13.2791 | -105.1864
13.2827 | -105.194
13.2864 | -105.2036
13.29 | -105.2148
13.2937 | -105.2265
13.2974 | -105.2412
13.301 | -105.2582
13.3047 | -105.2741
13.3083 | -105.2882
13.312 | -105.3064
13.3157 | -105.3236
13.3193 | -105.3394
13.323 | -105.3563
13.3266 | -105.3707
13.3303 | -105.3871
13.3339 | -105.4035
13.3376 | -105.4178
13.3413 | -105.4281
13.3449 | -105.4307
13.3486 | -105.4331
13.3522 | -105.4369
13.3559 | -105.435
13.3596 | -105.4283
13.3632 | -105.4174
13.3669 | -105.4019
13.3705 | -105.3857
13.3742 | -105.3709
13.3778 | -105.3497
13.3815 | -105.3178
13.3852 | -105.2826
13.3888 | -105.247
13.3925 | -105.2113
13.3961 | -105.1767
13.3998 | -105.1457
13.4034 | -105.1141
13.4071 | -105.0817
13.4108 | -105.0495
13.4144 | -105.0184
13.4181 | -104.9863
13.4217 | -104.96
13.4254 | -104.9333
13.4291 | -104.9065
13.4327 | -104.8837
13.4364 | -104.8599
13.44 | -104.8339
13.4437 | -104.8127
13.4473 | -104.7965
13.451 | -104.7818
13.4547 | -104.7696
13.4583 | -104.7609
13.462 | -104.7519
13.4656 | -104.7449
13.4693 | -104.7417
13.473 | -104.7406
13.4766 | -104.7399
13.4803 | -104.7409
13.4839 | -104.7445
13.4876 | -104.7508
13.4912 | -104.7603
13.4949 | -104.7708
13.4986 | -104.7826
13.5022 | -104.7946
13.5059 | -104.8044
13.5095 | -104.8145
13.5132 | -104.8281
13.5168 | -104.843
13.5205 | -104.8542
13.5242 | -104.8648
13.5278 | -104.8755
13.5315 | -104.8885
13.5351 | -104.9028
13.5388 | -104.9197
13.5425 | -104.9349
13.5461 | -104.9488
13.5498 | -104.9653
13.5534 | -104.9801
13.5571 | -104.9923
13.5607 | -105.0073
13.5644 | -105.0197
13.5681 | -105.0289
13.5717 | -105.0394
13.5754 | -105.0502
13.579 | -105.0591
13.5827 | -105.0672
13.5864 | -105.0708
13.59 | -105.073
13.5937 | -105.0732
13.5973 | -105.0749
13.601 | -105.0774
13.6046 | -105.0797
13.6083 | -105.0825
13.612 | -105.0849
13.6156 | -105.0859
13.6193 | -105.0875
13.6229 | -105.0885
13.6266 | -105.0885
13.6303 | -105.0933
13.6339 | -105.1007
13.6376 | -105.1072
13.6412 | -105.1134
13.6449 | -105.1185
13.6485 | -105.1225
13.6522 | -105.1285
13.6559 | -105.1398
13.6595 | -105.1506
13.6632 | -105.1617
13.6668 | -105.1733
13.6705 | -105.1871
13.6742 | -105.2012
13.6778 | -105.2132
13.6815 | -105.2224
13.6851 | -105.232
13.6888 | -105.244
13.6924 | -105.2541
13.6961 | -105.2665
13.6998 | -105.2806
13.7034 | -105.2938
13.7071 | -105.304
13.7107 | -105.3114
13.7144 | -105.3214
13.718 | -105.329
13.7217 | -105.337
13.7254 | -105.3403
13.729 | -105.3423
13.7327 | -105.3442
13.7363 | -105.3466
13.74 | -105.3466
13.7437 | -105.3462
13.7473 | -105.3463
13.751 | -105.3464
13.7546 | -105.3452
13.7583 | -105.34
13.7619 | -105.3381
13.7656 | -105.3349
13.7693 | -105.3277
13.7729 | -105.3237
13.7766 | -105.3184
13.7802 | -105.3107
13.7839 | -105.3034
13.7876 | -105.2961
13.7912 | -105.2902
13.7949 | -105.2831
13.7985 | -105.2761
13.8022 | -105.2697
13.8058 | -105.2632
13.8095 | -105.2578
13.8132 | -105.2542
13.8168 | -105.2525
13.8205 | -105.2546
13.8241 | -105.2588
13.8278 | -105.2596
13.8315 | -105.2612
13.8351 | -105.2641
13.8388 | -105.2651
13.8424 | -105.2682
13.8461 | -105.269
13.8497 | -105.27
13.8534 | -105.2702
13.8571 | -105.2693
13.8607 | -105.2703
13.8644 | -105.2723
13.868 | -105.2748
13.8717 | -105.2781
13.8753 | -105.2818
13.879 | -105.2834
13.8827 | -105.2864
13.8863 | -105.2877
13.89 | -105.2878
13.8936 | -105.2842
13.8973 | -105.2799
13.901 | -105.2759
13.9046 | -105.2725
13.9083 | -105.2691
13.9119 | -105.2655
13.9156 | -105.2598
13.9192 | -105.2522
13.9229 | -105.2447
13.9266 | -105.2401
13.9302 | -105.2365
13.9339 | -105.2302
13.9375 | -105.2259
13.9412 | -105.2222
13.9449 | -105.2161
13.9485 | -105.2077
13.9522 | -105.2032
13.9558 | -105.1956
13.9595 | -105.1726
13.9631 | -105.1054
13.9668 | -105.0533
13.9705 | -105.0166
13.9741 | -104.9769
13.9778 | -104.9333
13.9814 | -104.8861
13.9851 | -104.8327
13.9888 | -104.7743
13.9924 | -104.7109
13.9961 | -104.6425
13.9997 | -104.5707
14.0034 | -104.4941
14.007 | -104.4133
14.0107 | -104.3306
14.0144 | -104.2479
14.018 | -104.1669
14.0217 | -104.0862
14.0253 | -104.0073
14.029 | -103.9341
14.0326 | -103.8653
14.0363 | -103.8006
14.04 | -103.7448
14.0436 | -103.6967
14.0473 | -103.6583
14.0509 | -103.6282
14.0546 | -103.6069
14.0583 | -103.5964
14.0619 | -103.5952
14.0656 | -103.6043
14.0692 | -103.6236
14.0729 | -103.6518
14.0765 | -103.6902
14.0802 | -103.7414
14.0839 | -103.7987
14.0875 | -103.8625
14.0912 | -103.9352
14.0948 | -104.0139
14.0985 | -104.0978
14.1022 | -104.1872
14.1058 | -104.2791
14.1095 | -104.3713
14.1131 | -104.4612
14.1168 | -104.5533
14.1204 | -104.6441
14.1241 | -104.7337
14.1278 | -104.819
14.1314 | -104.8976
14.1351 | -104.9716
14.1387 | -105.0414
14.1424 | -105.1067
14.1461 | -105.1665
14.1497 | -105.2219
14.1534 | -105.2692
14.157 | -105.3135
14.1607 | -105.3644
14.1643 | -105.4428
14.168 | -105.4846
14.1717 | -105.4971
14.1753 | -105.5042
14.179 | -105.509
14.1826 | -105.5124
14.1863 | -105.5181
14.1899 | -105.5192
14.1936 | -105.5168
14.1973 | -105.5094
14.2009 | -105.4979
14.2046 | -105.4894
14.2082 | -105.4845
14.2119 | -105.4733
14.2156 | -105.457
14.2192 | -105.4422
14.2229 | -105.4263
14.2265 | -105.407
14.2302 | -105.3886
14.2338 | -105.371
14.2375 | -105.3547
14.2412 | -105.3355
14.2448 | -105.3124
14.2485 | -105.2856
14.2521 | -105.2643
14.2558 | -105.2409
14.2595 | -105.2179
14.2631 | -105.1963
14.2668 | -105.1756
14.2704 | -105.1544
14.2741 | -105.1356
14.2777 | -105.1196
14.2814 | -105.1036
14.2851 | -105.0887
14.2887 | -105.0746
14.2924 | -105.065
14.296 | -105.0603
14.2997 | -105.057
14.3034 | -105.052
14.307 | -105.0475
14.3107 | -105.0464
14.3143 | -105.0508
14.318 | -105.059
14.3216 | -105.0653
14.3253 | -105.0736
14.329 | -105.0831
14.3326 | -105.0931
14.3363 | -105.105
14.3399 | -105.1178
14.3436 | -105.1289
14.3473 | -105.1398
14.3509 | -105.1522
14.3546 | -105.1639
14.3582 | -105.1741
14.3619 | -105.1869
14.3655 | -105.1982
14.3692 | -105.2024
14.3729 | -105.2076
14.3765 | -105.2159
14.3802 | -105.2243
14.3838 | -105.2304
14.3875 | -105.2349
14.3911 | -105.2406
14.3948 | -105.2475
14.3985 | -105.2544
14.4021 | -105.2632
14.4058 | -105.273
14.4094 | -105.2797
14.4131 | -105.2839
14.4168 | -105.2911
14.4204 | -105.3006
14.4241 | -105.3066
14.4277 | -105.316
14.4314 | -105.3296
14.435 | -105.3426
14.4387 | -105.357
14.4424 | -105.3712
14.446 | -105.3883
14.4497 | -105.4091
14.4533 | -105.4309
14.457 | -105.4528
14.4607 | -105.4748
14.4643 | -105.4982
14.468 | -105.5238
14.4716 | -105.55
14.4753 | -105.5754
14.4789 | -105.6025
14.4826 | -105.6279
14.4863 | -105.6537
14.4899 | -105.6822
14.4936 | -105.7067
14.4972 | -105.7321
14.5009 | -105.7553
14.5046 | -105.7739
14.5082 | -105.7912
14.5119 | -105.8104
14.5155 | -105.8261
14.5192 | -105.8355
14.5228 | -105.8397
14.5265 | -105.8438
14.5302 | -105.8482
14.5338 | -105.8519
14.5375 | -105.8503
14.5411 | -105.8446
14.5448 | -105.8355
14.5484 | -105.826
14.5521 | -105.8128
14.5558 | -105.8001
14.5594 | -105.7862
14.5631 | -105.7691
14.5667 | -105.7497
14.5704 | -105.7271
14.5741 | -105.7045
14.5777 | -105.6835
14.5814 | -105.659
14.585 | -105.6344
14.5887 | -105.608
14.5923 | -105.5823
14.596 | -105.5591
14.5997 | -105.5329
14.6033 | -105.5052
14.607 | -105.4809
14.6106 | -105.4582
14.6143 | -105.4356
14.618 | -105.4124
14.6216 | -105.3914
14.6253 | -105.3709
14.6289 | -105.3492
14.6326 | -105.3303
14.6362 | -105.3107
14.6399 | -105.2921
14.6436 | -105.2753
14.6472 | -105.2594
14.6509 | -105.2433
14.6545 | -105.2273
14.6582 | -105.2137
14.6618 | -105.2012
14.6655 | -105.1892
14.6692 | -105.1786
14.6728 | -105.1703
14.6765 | -105.1616
14.6801 | -105.1517
14.6838 | -105.1446
14.6875 | -105.1389
14.6911 | -105.1363
14.6948 | -105.1343
14.6984 | -105.1287
14.7021 | -105.1252
14.7057 | -105.123
14.7094 | -105.1246
14.7131 | -105.1288
14.7167 | -105.1346
14.7204 | -105.1383
14.724 | -105.1426
14.7277 | -105.1507
14.7314 | -105.1592
14.735 | -105.1634
14.7387 | -105.1649
14.7423 | -105.1671
14.746 | -105.174
14.7496 | -105.1791
14.7533 | -105.1824
14.757 | -105.1871
14.7606 | -105.1915
14.7643 | -105.1935
14.7679 | -105.1957
14.7716 | -105.1995
14.7753 | -105.1993
14.7789 | -105.203
14.7826 | -105.2034
14.7862 | -105.2029
14.7899 | -105.2027
14.7935 | -105.1991
14.7972 | -105.19
14.8009 | -105.1844
14.8045 | -105.1803
14.8082 | -105.1741
14.8118 | -105.1664
14.8155 | -105.1587
14.8192 | -105.1483
14.8228 | -105.138
14.8265 | -105.1254
14.8301 | -105.1152
14.8338 | -105.1036
14.8374 | -105.0933
14.8411 | -105.082
14.8448 | -105.0715
14.8484 | -105.0634
14.8521 | -105.0546
14.8557 | -105.0444
14.8594 | -105.0346
14.863 | -105.0292
14.8667 | -105.0237
14.8704 | -105.0186
14.874 | -105.0157
14.8777 | -105.0157
14.8813 | -105.0192
14.885 | -105.0227
14.8887 | -105.027
14.8923 | -105.0341
14.896 | -105.039
14.8996 | -105.0452
14.9033 | -105.0535
14.9069 | -105.064
14.9106 | -105.0811
14.9143 | -105.0973
14.9179 | -105.1118
14.9216 | -105.126
14.9252 | -105.1405
14.9289 | -105.1552
14.9326 | -105.1704
14.9362 | -105.1867
14.9399 | -105.2019
14.9435 | -105.216
14.9472 | -105.2325
14.9508 | -105.2504
14.9545 | -105.2643
14.9582 | -105.2753
14.9618 | -105.2854
14.9655 | -105.2938
14.9691 | -105.2975
14.9728 | -105.3013
14.9765 | -105.3059
14.9801 | -105.3062
14.9838 | -105.3005
14.9874 | -105.2831
14.9911 | -105.2654
14.9947 | -105.2526
14.9984 | -105.2399
15.0021 | -105.2248
15.0057 | -105.2072
15.0094 | -105.1904
15.013 | -105.174
15.0167 | -105.1539
15.0203 | -105.1305
15.024 | -105.1038
15.0277 | -105.0781
15.0313 | -105.0537
15.035 | -105.0303
15.0386 | -105.0065
15.0423 | -104.9826
15.046 | -104.9567
15.0496 | -104.9345
15.0533 | -104.9135
15.0569 | -104.8952
15.0606 | -104.8769
15.0642 | -104.8587
15.0679 | -104.8436
15.0716 | -104.8326
15.0752 | -104.8237
15.0789 | -104.8161
15.0825 | -104.8118
15.0862 | -104.8098
15.0899 | -104.8089
15.0935 | -104.8081
15.0972 | -104.8074
15.1008 | -104.813
15.1045 | -104.8206
15.1081 | -104.832
15.1118 | -104.8469
15.1155 | -104.8631
15.1191 | -104.8762
15.1228 | -104.8898
15.1264 | -104.9051
15.1301 | -104.9224
15.1338 | -104.9407
15.1374 | -104.9569
15.1411 | -104.9723
15.1447 | -104.9873
15.1484 | -105.0052
15.152 | -105.0214
15.1557 | -105.0333
15.1594 | -105.0428
15.163 | -105.0468
15.1667 | -105.0485
15.1703 | -105.0506
15.174 | -105.0517
15.1776 | -105.0489
15.1813 | -105.0441
15.185 | -105.0393
15.1886 | -105.0335
15.1923 | -105.0227
15.1959 | -105.0136
15.1996 | -105.0066
15.2033 | -105.007
15.2069 | -105.0016
15.2106 | -104.986
15.2142 | -104.9684
15.2179 | -104.9519
15.2215 | -104.9355
15.2252 | -104.9217
15.2289 | -104.9068
15.2325 | -104.8904
15.2362 | -104.8751
15.2398 | -104.8621
15.2435 | -104.8498
15.2472 | -104.836
15.2508 | -104.8201
15.2545 | -104.8071
15.2581 | -104.7975
15.2618 | -104.7884
15.2654 | -104.7774
15.2691 | -104.7667
15.2728 | -104.7567
15.2764 | -104.749
15.2801 | -104.745
15.2837 | -104.7403
15.2874 | -104.7355
15.2911 | -104.7314
15.2947 | -104.7303
15.2984 | -104.7306
15.302 | -104.7357
15.3057 | -104.7382
15.3093 | -104.7431
15.313 | -104.7495
15.3167 | -104.7542
15.3203 | -104.7553
15.324 | -104.7594
15.3276 | -104.7648
15.3313 | -104.7703
15.3349 | -104.7754
15.3386 | -104.7819
15.3423 | -104.7867
15.3459 | -104.7904
15.3496 | -104.7952
15.3532 | -104.7996
15.3569 | -104.8014
15.3606 | -104.8042
15.3642 | -104.8068
15.3679 | -104.8042
15.3715 | -104.8002
15.3752 | -104.798
15.3788 | -104.7947
15.3825 | -104.7913
15.3862 | -104.7871
15.3898 | -104.7836
15.3935 | -104.778
15.3971 | -104.7743
15.4008 | -104.7699
15.4045 | -104.7632
15.4081 | -104.7546
15.4118 | -104.749
15.4154 | -104.7428
15.4191 | -104.7374
15.4227 | -104.7285
15.4264 | -104.7205
15.4301 | -104.7174
15.4337 | -104.7128
15.4374 | -104.7071
15.441 | -104.7013
15.4447 | -104.6945
15.4484 | -104.6894
15.452 | -104.6856
15.4557 | -104.6806
15.4593 | -104.6769
15.463 | -104.6719
15.4666 | -104.6666
15.4703 | -104.6629
15.474 | -104.657
15.4776 | -104.6523
15.4813 | -104.6473
15.4849 | -104.6412
15.4886 | -104.6375
15.4922 | -104.635
15.4959 | -104.6321
15.4996 | -104.6285
15.5032 | -104.6235
15.5069 | -104.6192
15.5105 | -104.616
15.5142 | -104.6106
15.5179 | -104.6039
15.5215 | -104.5977
15.5252 | -104.5909
15.5288 | -104.5851
15.5325 | -104.5798
15.5361 | -104.5746
15.5398 | -104.5715
15.5435 | -104.571
15.5471 | -104.5687
15.5508 | -104.5631
15.5544 | -104.5585
15.5581 | -104.5541
15.5618 | -104.5527
15.5654 | -104.5509
15.5691 | -104.5453
15.5727 | -104.5385
15.5764 | -104.5333
15.58 | -104.5323
15.5837 | -104.5308
15.5874 | -104.5308
15.591 | -104.5325
15.5947 | -104.5333
15.5983 | -104.5354
15.602 | -104.5392
15.6057 | -104.5441
15.6093 | -104.5481
15.613 | -104.5497
15.6166 | -104.5524
15.6203 | -104.5572
15.6239 | -104.5629
15.6276 | -104.5678
15.6313 | -104.5746
15.6349 | -104.5802
15.6386 | -104.5881
15.6422 | -104.5973
15.6459 | -104.6081
15.6495 | -104.6175
15.6532 | -104.6251
15.6569 | -104.6354
15.6605 | -104.6463
15.6642 | -104.6562
15.6678 | -104.6664
15.6715 | -104.6753
15.6752 | -104.6849
15.6788 | -104.6966
15.6825 | -104.7052
15.6861 | -104.7137
15.6898 | -104.7255
15.6934 | -104.7352
15.6971 | -104.745
15.7008 | -104.7516
15.7044 | -104.7593
15.7081 | -104.7659
15.7117 | -104.7723
15.7154 | -104.7824
15.7191 | -104.793
15.7227 | -104.8013
15.7264 | -104.8119
15.73 | -104.8196
15.7337 | -104.8253
15.7373 | -104.8304
15.741 | -104.8349
15.7447 | -104.8405
15.7483 | -104.8442
15.752 | -104.8428
15.7556 | -104.84
15.7593 | -104.8392
15.763 | -104.8361
15.7666 | -104.8283
15.7703 | -104.8233
15.7739 | -104.8206
15.7776 | -104.8137
15.7812 | -104.8053
15.7849 | -104.7961
15.7886 | -104.7882
15.7922 | -104.7826
15.7959 | -104.7764
15.7995 | -104.77
15.8032 | -104.7613
15.8069 | -104.7518
15.8105 | -104.7416
15.8142 | -104.7322
15.8178 | -104.7231
15.8215 | -104.7146
15.8251 | -104.7084
15.8288 | -104.7011
15.8325 | -104.7021
15.8361 | -104.6953
15.8398 | -104.6919
15.8434 | -104.6898
15.8471 | -104.6872
15.8507 | -104.6856
15.8544 | -104.6873
15.8581 | -104.6876
15.8617 | -104.6915
15.8654 | -104.6948
15.869 | -104.6985
15.8727 | -104.703
15.8764 | -104.7093
15.88 | -104.716
15.8837 | -104.7234
15.8873 | -104.7309
15.891 | -104.7382
15.8946 | -104.7442
15.8983 | -104.7533
15.902 | -104.762
15.9056 | -104.7708
15.9093 | -104.7805
15.9129 | -104.7882
15.9166 | -104.7932
15.9203 | -104.8001
15.9239 | -104.8093
15.9276 | -104.8204
15.9312 | -104.8285
15.9349 | -104.8349
15.9385 | -104.8393
15.9422 | -104.8385
15.9459 | -104.8367
15.9495 | -104.8388
15.9532 | -104.839
15.9568 | -104.8408
15.9605 | -104.8432
15.9642 | -104.8444
15.9678 | -104.8451
15.9715 | -104.8445
15.9751 | -104.8436
15.9788 | -104.8429
15.9824 | -104.8391
15.9861 | -104.8338
15.9898 | -104.8315
15.9934 | -104.8319
15.9971 | -104.8293
16.0007 | -104.8236
16.0044 | -104.819
16.008 | -104.8151
16.0117 | -104.8096
16.0154 | -104.8036
16.019 | -104.7978
16.0227 | -104.7931
16.0263 | -104.7831
16.03 | -104.7728
16.0337 | -104.7663
16.0373 | -104.7615
16.041 | -104.756
16.0446 | -104.7521
16.0483 | -104.7486
16.0519 | -104.743
16.0556 | -104.7401
16.0593 | -104.7313
16.0629 | -104.7253
16.0666 | -104.7241
16.0702 | -104.7196
16.0739 | -104.7146
16.0776 | -104.7113
16.0812 | -104.7094
16.0849 | -104.7069
16.0885 | -104.7064
16.0922 | -104.7021
16.0958 | -104.6986
16.0995 | -104.697
16.1032 | -104.6989
16.1068 | -104.6999
16.1105 | -104.7007
16.1141 | -104.7016
16.1178 | -104.705
16.1214 | -104.7086
16.1251 | -104.7131
16.1288 | -104.7186
16.1324 | -104.7238
16.1361 | -104.7273
16.1397 | -104.7312
16.1434 | -104.7387
16.1471 | -104.7483
16.1507 | -104.7573
16.1544 | -104.7611
16.158 | -104.7681
16.1617 | -104.7796
16.1653 | -104.7936
16.169 | -104.8062
16.1727 | -104.8197
16.1763 | -104.8342
16.18 | -104.847
16.1836 | -104.8612
16.1873 | -104.8735
16.191 | -104.8883
16.1946 | -104.9041
16.1983 | -104.9213
16.2019 | -104.9373
16.2056 | -104.953
16.2092 | -104.9677
16.2129 | -104.985
16.2166 | -105.0012
16.2202 | -105.0156
16.2239 | -105.0278
16.2275 | -105.04
16.2312 | -105.0519
16.2349 | -105.0652
16.2385 | -105.0793
16.2422 | -105.0923
16.2458 | -105.1016
16.2495 | -105.1078
16.2531 | -105.1143
16.2568 | -105.1257
16.2605 | -105.1372
16.2641 | -105.1411
16.2678 | -105.1419
16.2714 | -105.1434
16.2751 | -105.1445
16.2788 | -105.1441
16.2824 | -105.1454
16.2861 | -105.1447
16.2897 | -105.1454
16.2934 | -105.1423
16.297 | -105.1343
16.3007 | -105.1283
16.3044 | -105.1226
16.308 | -105.1129
16.3117 | -105.1016
16.3153 | -105.0897
16.319 | -105.0776
16.3226 | -105.0684
16.3263 | -105.0569
16.33 | -105.0427
16.3336 | -105.0278
16.3373 | -105.0144
16.3409 | -104.9998
16.3446 | -104.9864
16.3483 | -104.9712
16.3519 | -104.9591
16.3556 | -104.9471
16.3592 | -104.9421
16.3629 | -104.9307
16.3665 | -104.9223
16.3702 | -104.9189
16.3739 | -104.9158
16.3775 | -104.9122
16.3812 | -104.9069
16.3848 | -104.9045
16.3885 | -104.9067
16.3922 | -104.9074
16.3958 | -104.9083
16.3995 | -104.9114
16.4031 | -104.9157
16.4068 | -104.9222
16.4104 | -104.9302
16.4141 | -104.9401
16.4178 | -104.9499
16.4214 | -104.9616
16.4251 | -104.9715
16.4287 | -104.9808
16.4324 | -104.9915
16.4361 | -105.004
16.4397 | -105.0172
16.4434 | -105.0323
16.447 | -105.0462
16.4507 | -105.0614
16.4543 | -105.0752
16.458 | -105.0897
16.4617 | -105.1045
16.4653 | -105.1212
16.469 | -105.1365
16.4726 | -105.1478
16.4763 | -105.1603
16.4799 | -105.1765
16.4836 | -105.1915
16.4873 | -105.2047
16.4909 | -105.2132
16.4946 | -105.2209
16.4982 | -105.233
16.5019 | -105.2454
16.5056 | -105.2555
16.5092 | -105.264
16.5129 | -105.2736
16.5165 | -105.2828
16.5202 | -105.2904
16.5238 | -105.301
16.5275 | -105.3087
16.5312 | -105.3147
16.5348 | -105.3184
16.5385 | -105.3237
16.5421 | -105.3279
16.5458 | -105.3277
16.5495 | -105.3262
16.5531 | -105.323
16.5568 | -105.3151
16.5604 | -105.3085
16.5641 | -105.3023
16.5677 | -105.2903
16.5714 | -105.276
16.5751 | -105.2598
16.5787 | -105.2421
16.5824 | -105.2243
16.586 | -105.2044
16.5897 | -105.1858
16.5934 | -105.1652
16.597 | -105.1441
16.6007 | -105.1242
16.6043 | -105.1003
16.608 | -105.0774
16.6116 | -105.0572
16.6153 | -105.0386
16.619 | -105.0227
16.6226 | -105.0065
16.6263 | -104.9902
16.6299 | -104.9752
16.6336 | -104.9616
16.6372 | -104.9486
16.6409 | -104.9373
16.6446 | -104.9277
16.6482 | -104.9179
16.6519 | -104.9088
16.6555 | -104.9039
16.6592 | -104.8991
16.6629 | -104.8958
16.6665 | -104.8955
16.6702 | -104.8959
16.6738 | -104.8969
16.6775 | -104.8982
16.6811 | -104.9002
16.6848 | -104.905
16.6885 | -104.9087
16.6921 | -104.9134
16.6958 | -104.9207
16.6994 | -104.9263
16.7031 | -104.9314
16.7068 | -104.9362
16.7104 | -104.9441
16.7141 | -104.9521
16.7177 | -104.9591
16.7214 | -104.9668
16.725 | -104.9765
16.7287 | -104.9882
16.7324 | -104.9971
16.736 | -105.0021
16.7397 | -105.0074
16.7433 | -105.0128
16.747 | -105.02
16.7507 | -105.0253
16.7543 | -105.028
16.758 | -105.034
16.7616 | -105.0392
16.7653 | -105.044
16.7689 | -105.0496
16.7726 | -105.0546
16.7763 | -105.0556
16.7799 | -105.0582
16.7836 | -105.062
16.7872 | -105.0656
16.7909 | -105.0671
16.7945 | -105.0683
16.7982 | -105.069
16.8019 | -105.0707
16.8055 | -105.0745
16.8092 | -105.0798
16.8128 | -105.0879
16.8165 | -105.0937
16.8202 | -105.0963
16.8238 | -105.0985
16.8275 | -105.1001
16.8311 | -105.1039
16.8348 | -105.1059
16.8384 | -105.1074
16.8421 | -105.1136
16.8458 | -105.1183
16.8494 | -105.1206
16.8531 | -105.1221
16.8567 | -105.1245
16.8604 | -105.1274
16.8641 | -105.1322
16.8677 | -105.1362
16.8714 | -105.1408
16.875 | -105.1456
16.8787 | -105.1514
16.8823 | -105.1552
16.886 | -105.1593
16.8897 | -105.165
16.8933 | -105.1699
16.897 | -105.1746
16.9006 | -105.1803
16.9043 | -105.1847
16.908 | -105.1873
16.9116 | -105.1887
16.9153 | -105.1911
16.9189 | -105.1971
16.9226 | -105.2031
16.9262 | -105.2071
16.9299 | -105.211
16.9336 | -105.2141
16.9372 | -105.2175
16.9409 | -105.2226
16.9445 | -105.2248
16.9482 | -105.2299
16.9519 | -105.2342
16.9555 | -105.2376
16.9592 | -105.2425
16.9628 | -105.2464
16.9665 | -105.2456
16.9701 | -105.2449
16.9738 | -105.2456
16.9775 | -105.2491
16.9811 | -105.252
16.9848 | -105.2543
16.9884 | -105.2524
16.9921 | -105.2497
16.9957 | -105.2488
16.9994 | -105.247
17.0031 | -105.2449
17.0067 | -105.243
17.0104 | -105.2402
17.014 | -105.2359
17.0177 | -105.2327
17.0214 | -105.229
17.025 | -105.224
17.0287 | -105.2185
17.0323 | -105.2144
17.036 | -105.2108
17.0396 | -105.2076
17.0433 | -105.2002
17.047 | -105.1947
17.0506 | -105.1893
17.0543 | -105.1834
17.0579 | -105.1804
17.0616 | -105.1792
17.0653 | -105.177
17.0689 | -105.1763
17.0726 | -105.1764
17.0762 | -105.1781
17.0799 | -105.1796
17.0835 | -105.1789
17.0872 | -105.1774
17.0909 | -105.1781
17.0945 | -105.1805
17.0982 | -105.1819
17.1018 | -105.1836
17.1055 | -105.1858
17.1091 | -105.1882
17.1128 | -105.1913
17.1165 | -105.1957
17.1201 | -105.1997
17.1238 | -105.2048
17.1274 | -105.2109
17.1311 | -105.2157
17.1348 | -105.2202
17.1384 | -105.2251
17.1421 | -105.2292
17.1457 | -105.2349
17.1494 | -105.2417
17.153 | -105.2505
17.1567 | -105.2593
17.1604 | -105.2685
17.164 | -105.2774
17.1677 | -105.2863
17.1713 | -105.297
17.175 | -105.3071
17.1787 | -105.3181
17.1823 | -105.3254
17.186 | -105.3326
17.1896 | -105.3437
17.1933 | -105.3519
17.1969 | -105.3585
17.2006 | -105.366
17.2043 | -105.3712
17.2079 | -105.3764
17.2116 | -105.3835
17.2152 | -105.3904
17.2189 | -105.3954
17.2226 | -105.3976
17.2262 | -105.3988
17.2299 | -105.4029
17.2335 | -105.4058
17.2372 | -105.4099
17.2408 | -105.4129
17.2445 | -105.4129
17.2482 | -105.4151
17.2518 | -105.418
17.2555 | -105.4196
17.2591 | -105.4208
17.2628 | -105.4241
17.2664 | -105.4271
17.2701 | -105.4276
17.2738 | -105.4278
17.2774 | -105.428
17.2811 | -105.4256
17.2847 | -105.42
17.2884 | -105.4199
17.2921 | -105.4189
17.2957 | -105.4173
17.2994 | -105.4148
17.303 | -105.4101
17.3067 | -105.4089
17.3103 | -105.4098
17.314 | -105.4076
17.3177 | -105.4052
17.3213 | -105.4025
17.325 | -105.4028
17.3286 | -105.4036
17.3323 | -105.4063
17.336 | -105.4111
17.3396 | -105.415
17.3433 | -105.418
17.3469 | -105.421
17.3506 | -105.4256
17.3542 | -105.4313
17.3579 | -105.4355
17.3616 | -105.4395
17.3652 | -105.4454
17.3689 | -105.4496
17.3725 | -105.4544
17.3762 | -105.4608
17.3799 | -105.466
17.3835 | -105.4719
17.3872 | -105.4779
17.3908 | -105.4822
17.3945 | -105.4834
17.3981 | -105.4867
17.4018 | -105.4898
17.4055 | -105.4881
17.4091 | -105.4841
17.4128 | -105.474
17.4164 | -105.4689
17.4201 | -105.4632
17.4238 | -105.4585
17.4274 | -105.4549
17.4311 | -105.451
17.4347 | -105.4454
17.4384 | -105.4362
17.442 | -105.4281
17.4457 | -105.4201
17.4494 | -105.4116
17.453 | -105.4028
17.4567 | -105.3947
17.4603 | -105.3845
17.464 | -105.3728
17.4676 | -105.3635
17.4713 | -105.3551
17.475 | -105.3463
17.4786 | -105.3365
17.4823 | -105.328
17.4859 | -105.317
17.4896 | -105.3092
17.4933 | -105.3052
17.4969 | -105.3013
17.5006 | -105.299
17.5042 | -105.2985
17.5079 | -105.2989
17.5115 | -105.2978
17.5152 | -105.2991
17.5189 | -105.3052
17.5225 | -105.3125
17.5262 | -105.3192
17.5298 | -105.3287
17.5335 | -105.3376
17.5372 | -105.3453
17.5408 | -105.3541
17.5445 | -105.3633
17.5481 | -105.3729
17.5518 | -105.3848
17.5554 | -105.3946
17.5591 | -105.4045
17.5628 | -105.4144
17.5664 | -105.4255
17.5701 | -105.4389
17.5737 | -105.4482
17.5774 | -105.458
17.5811 | -105.467
17.5847 | -105.4749
17.5884 | -105.4812
17.592 | -105.4873
17.5957 | -105.4953
17.5993 | -105.5006
17.603 | -105.5039
17.6067 | -105.5078
17.6103 | -105.5108
17.614 | -105.5111
17.6176 | -105.5128
17.6213 | -105.5133
17.6249 | -105.5127
17.6286 | -105.5125
17.6323 | -105.5115
17.6359 | -105.5081
17.6396 | -105.5047
17.6432 | -105.5021
17.6469 | -105.4946
17.6506 | -105.4871
17.6542 | -105.4835
17.6579 | -105.4795
17.6615 | -105.4757
17.6652 | -105.4733
17.6688 | -105.47
17.6725 | -105.4638
17.6762 | -105.4571
17.6798 | -105.4496
17.6835 | -105.443
17.6871 | -105.4354
17.6908 | -105.4309
17.6945 | -105.424
17.6981 | -105.4154
17.7018 | -105.4057
17.7054 | -105.3965
17.7091 | -105.3887
17.7127 | -105.3828
17.7164 | -105.3761
17.7201 | -105.3677
17.7237 | -105.3591
17.7274 | -105.3507
17.731 | -105.3413
17.7347 | -105.3333
17.7384 | -105.3241
17.742 | -105.3127
17.7457 | -105.3013
17.7493 | -105.2864
17.753 | -105.2739
17.7566 | -105.2612
17.7603 | -105.2481
17.764 | -105.2349
17.7676 | -105.2223
17.7713 | -105.2083
17.7749 | -105.1923
17.7786 | -105.1803
17.7822 | -105.1702
17.7859 | -105.162
17.7896 | -105.1517
17.7932 | -105.1399
17.7969 | -105.1282
17.8005 | -105.1163
17.8042 | -105.1039
17.8079 | -105.0941
17.8115 | -105.0838
17.8152 | -105.0758
17.8188 | -105.0651
17.8225 | -105.0535
17.8261 | -105.047
17.8298 | -105.0394
17.8335 | -105.0319
17.8371 | -105.0238
17.8408 | -105.018
17.8444 | -105.0138
17.8481 | -105.009
17.8518 | -105.0061
17.8554 | -105.0048
17.8591 | -105.0045
17.8627 | -105.0052
17.8664 | -105.0068
17.87 | -105.0092
17.8737 | -105.0112
17.8774 | -105.0141
17.881 | -105.0178
17.8847 | -105.0197
17.8883 | -105.0245
17.892 | -105.0314
17.8957 | -105.039
17.8993 | -105.0479
17.903 | -105.0578
17.9066 | -105.066
17.9103 | -105.0752
17.9139 | -105.0836
17.9176 | -105.0923
17.9213 | -105.1028
17.9249 | -105.1138
17.9286 | -105.1271
17.9322 | -105.139
17.9359 | -105.151
17.9395 | -105.1662
17.9432 | -105.1798
17.9469 | -105.1901
17.9505 | -105.2029
17.9542 | -105.2162
17.9578 | -105.2278
17.9615 | -105.2401
17.9652 | -105.2525
17.9688 | -105.2638
17.9725 | -105.2756
17.9761 | -105.2877
17.9798 | -105.2986
17.9834 | -105.3106
17.9871 | -105.3218
17.9908 | -105.3302
17.9944 | -105.3416
17.9981 | -105.354
18.0017 | -105.3659
18.0054 | -105.3789
18.0091 | -105.3904
18.0127 | -105.4026
18.0164 | -105.4156
18.02 | -105.4288
18.0237 | -105.4428
18.0273 | -105.4576
18.031 | -105.4728
18.0347 | -105.4827
18.0383 | -105.4917
18.042 | -105.5002
18.0456 | -105.5113
18.0493 | -105.5233
18.053 | -105.5353
18.0566 | -105.5466
18.0603 | -105.5575
18.0639 | -105.5687
18.0676 | -105.5797
18.0712 | -105.5886
18.0749 | -105.6002
18.0786 | -105.6092
18.0822 | -105.6152
18.0859 | -105.6234
18.0895 | -105.63
18.0932 | -105.6361
18.0968 | -105.6425
18.1005 | -105.6471
18.1042 | -105.6482
18.1078 | -105.6491
18.1115 | -105.6525
18.1151 | -105.6551
18.1188 | -105.6558
18.1225 | -105.6559
18.1261 | -105.6548
18.1298 | -105.6534
18.1334 | -105.653
18.1371 | -105.6498
18.1407 | -105.6507
18.1444 | -105.6495
18.1481 | -105.6447
18.1517 | -105.6415
18.1554 | -105.6377
18.159 | -105.6317
18.1627 | -105.6255
18.1664 | -105.6203
18.17 | -105.6148
18.1737 | -105.6054
18.1773 | -105.5937
18.181 | -105.5826
18.1846 | -105.569
18.1883 | -105.5573
18.192 | -105.5434
18.1956 | -105.5261
18.1993 | -105.5083
18.2029 | -105.489
18.2066 | -105.471
18.2103 | -105.4517
18.2139 | -105.4315
18.2176 | -105.4116
18.2212 | -105.3889
18.2249 | -105.362
18.2285 | -105.3352
18.2322 | -105.3095
18.2359 | -105.2815
18.2395 | -105.2504
18.2432 | -105.217
18.2468 | -105.1865
18.2505 | -105.1573
18.2541 | -105.1268
18.2578 | -105.0955
18.2615 | -105.0635
18.2651 | -105.0326
18.2688 | -105.0007
18.2724 | -104.9683
18.2761 | -104.9365
18.2798 | -104.9037
18.2834 | -104.8709
18.2871 | -104.8393
18.2907 | -104.8092
18.2944 | -104.7811
18.298 | -104.7532
18.3017 | -104.7273
18.3054 | -104.7022
18.309 | -104.6776
18.3127 | -104.6554
18.3163 | -104.6351
18.32 | -104.6156
18.3237 | -104.5967
18.3273 | -104.5785
18.331 | -104.5619
18.3346 | -104.5468
18.3383 | -104.5365
18.3419 | -104.5284
18.3456 | -104.5207
18.3493 | -104.5153
18.3529 | -104.5128
18.3566 | -104.5103
18.3602 | -104.5099
18.3639 | -104.5147
18.3676 | -104.5208
18.3712 | -104.5279
18.3749 | -104.535
18.3785 | -104.5449
18.3822 | -104.5583
18.3858 | -104.5731
18.3895 | -104.5897
18.3932 | -104.6072
18.3968 | -104.6285
18.4005 | -104.6502
18.4041 | -104.673
18.4078 | -104.6973
18.4114 | -104.7222
18.4151 | -104.7502
18.4188 | -104.7804
18.4224 | -104.8097
18.4261 | -104.8395
18.4297 | -104.8712
18.4334 | -104.9031
18.4371 | -104.9366
18.4407 | -104.9702
18.4444 | -105.0024
18.448 | -105.0329
18.4517 | -105.064
18.4553 | -105.0951
18.459 | -105.1269
18.4627 | -105.1587
18.4663 | -105.1879
18.47 | -105.2066
18.4736 | -105.2303
18.4773 | -105.2527
18.481 | -105.2725
18.4846 | -105.2911
18.4883 | -105.3106
18.4919 | -105.3258
18.4956 | -105.3363
18.4992 | -105.347
18.5029 | -105.3595
18.5066 | -105.3697
18.5102 | -105.3766
18.5139 | -105.3818
18.5175 | -105.3826
18.5212 | -105.3809
18.5249 | -105.3776
18.5285 | -105.374
18.5322 | -105.369
18.5358 | -105.3636
18.5395 | -105.3557
18.5431 | -105.3462
18.5468 | -105.3349
18.5505 | -105.3206
18.5541 | -105.3064
18.5578 | -105.2917
18.5614 | -105.2734
18.5651 | -105.2543
18.5688 | -105.2349
18.5724 | -105.2139
18.5761 | -105.1934
18.5797 | -105.176
18.5834 | -105.1579
18.587 | -105.1395
18.5907 | -105.1233
18.5944 | -105.1071
18.598 | -105.0903
18.6017 | -105.0742
18.6053 | -105.059
18.609 | -105.0451
18.6126 | -105.0326
18.6163 | -105.0202
18.62 | -105.0094
18.6236 | -105.0035
18.6273 | -104.9964
18.6309 | -104.9902
18.6346 | -104.9838
18.6383 | -104.9814
18.6419 | -104.9824
18.6456 | -104.9823
18.6492 | -104.9805
18.6529 | -104.9823
18.6565 | -104.9854
18.6602 | -104.9884
18.6639 | -104.9921
18.6675 | -104.9971
18.6712 | -105.0046
18.6748 | -105.0123
18.6785 | -105.0192
18.6822 | -105.0265
18.6858 | -105.0366
18.6895 | -105.0461
18.6931 | -105.0557
18.6968 | -105.0677
18.7004 | -105.0797
18.7041 | -105.0913
18.7078 | -105.1017
18.7114 | -105.1128
18.7151 | -105.1238
18.7187 | -105.1363
18.7224 | -105.1486
18.7261 | -105.1604
18.7297 | -105.1721
18.7334 | -105.1855
18.737 | -105.1988
18.7407 | -105.2121
18.7443 | -105.2258
18.748 | -105.2411
18.7517 | -105.2556
18.7553 | -105.2671
18.759 | -105.2784
18.7626 | -105.2908
18.7663 | -105.3019
18.7699 | -105.3081
18.7736 | -105.3132
18.7773 | -105.3174
18.7809 | -105.3228
18.7846 | -105.3295
18.7882 | -105.3352
18.7919 | -105.3401
18.7956 | -105.342
18.7992 | -105.3446
18.8029 | -105.3483
18.8065 | -105.3501
18.8102 | -105.3513
18.8138 | -105.352
18.8175 | -105.3536
18.8212 | -105.353
18.8248 | -105.3512
18.8285 | -105.3527
18.8321 | -105.3524
18.8358 | -105.3521
18.8395 | -105.3538
18.8431 | -105.354
18.8468 | -105.3517
18.8504 | -105.3496
18.8541 | -105.3474
18.8577 | -105.3467
18.8614 | -105.3479
18.8651 | -105.3494
18.8687 | -105.3511
18.8724 | -105.3529
18.876 | -105.3552
18.8797 | -105.3595
18.8834 | -105.3637
18.887 | -105.3697
18.8907 | -105.3724
18.8943 | -105.3754
18.898 | -105.3801
18.9016 | -105.3857
18.9053 | -105.3891
18.909 | -105.3927
18.9126 | -105.3999
18.9163 | -105.4099
18.9199 | -105.4167
18.9236 | -105.4216
18.9272 | -105.4283
18.9309 | -105.435
18.9346 | -105.4419
18.9382 | -105.4456
18.9419 | -105.4526
18.9455 | -105.4607
18.9492 | -105.4692
18.9529 | -105.4768
18.9565 | -105.4845
18.9602 | -105.4927
18.9638 | -105.4989
18.9675 | -105.5033
18.9711 | -105.5078
18.9748 | -105.513
18.9785 | -105.5153
18.9821 | -105.5165
18.9858 | -105.5184
18.9894 | -105.5192
18.9931 | -105.5194
18.9968 | -105.5137
19.0004 | -105.5125
19.0041 | -105.5091
19.0077 | -105.5052
19.0114 | -105.502
19.015 | -105.4978
19.0187 | -105.4918
19.0224 | -105.4863
19.026 | -105.4802
19.0297 | -105.4732
19.0333 | -105.4653
19.037 | -105.4543
19.0407 | -105.4469
19.0443 | -105.4364
19.048 | -105.4251
19.0516 | -105.4122
19.0553 | -105.3971
19.0589 | -105.3845
19.0626 | -105.3725
19.0663 | -105.3587
19.0699 | -105.3472
19.0736 | -105.3353
19.0772 | -105.3251
19.0809 | -105.3156
19.0845 | -105.3064
19.0882 | -105.2973
19.0919 | -105.2885
19.0955 | -105.2799
19.0992 | -105.2714
19.1028 | -105.2643
19.1065 | -105.2558
19.1102 | -105.2497
19.1138 | -105.245
19.1175 | -105.2425
19.1211 | -105.241
19.1248 | -105.2392
19.1284 | -105.2383
19.1321 | -105.2366
19.1358 | -105.2361
19.1394 | -105.2335
19.1431 | -105.2329
19.1467 | -105.235
19.1504 | -105.2367
19.1541 | -105.2409
19.1577 | -105.2448
19.1614 | -105.2488
19.165 | -105.2536
19.1687 | -105.2577
19.1723 | -105.2637
19.176 | -105.2699
19.1797 | -105.2754
19.1833 | -105.2818
19.187 | -105.2876
19.1906 | -105.2926
19.1943 | -105.2997
19.198 | -105.3075
19.2016 | -105.3142
19.2053 | -105.3196
19.2089 | -105.3265
19.2126 | -105.3296
19.2162 | -105.3295
19.2199 | -105.3308
19.2236 | -105.3341
19.2272 | -105.3366
19.2309 | -105.3384
19.2345 | -105.3413
19.2382 | -105.3441
19.2418 | -105.3462
19.2455 | -105.3475
19.2492 | -105.348
19.2528 | -105.3476
19.2565 | -105.3435
19.2601 | -105.3392
19.2638 | -105.3373
19.2675 | -105.3357
19.2711 | -105.3348
19.2748 | -105.3336
19.2784 | -105.3306
19.2821 | -105.3264
19.2857 | -105.3217
19.2894 | -105.3192
19.2931 | -105.3169
19.2967 | -105.3168
19.3004 | -105.3143
19.304 | -105.3105
19.3077 | -105.309
19.3114 | -105.3096
19.315 | -105.3084
19.3187 | -105.3075
19.3223 | -105.305
19.326 | -105.3015
19.3296 | -105.2997
19.3333 | -105.2955
19.337 | -105.2906
19.3406 | -105.2884
19.3443 | -105.2842
19.3479 | -105.2792
19.3516 | -105.2744
19.3553 | -105.2676
19.3589 | -105.2589
19.3626 | -105.2495
19.3662 | -105.242
19.3699 | -105.2362
19.3735 | -105.2297
19.3772 | -105.2222
19.3809 | -105.2144
19.3845 | -105.2051
19.3882 | -105.1951
19.3918 | -105.1842
19.3955 | -105.173
19.3991 | -105.162
19.4028 | -105.1504
19.4065 | -105.1394
19.4101 | -105.1274
19.4138 | -105.1168
19.4174 | -105.1043
19.4211 | -105.0902
19.4248 | -105.0767
19.4284 | -105.0619
19.4321 | -105.0463
19.4357 | -105.0321
19.4394 | -105.0193
19.443 | -105.0057
19.4467 | -104.9919
19.4504 | -104.9794
19.454 | -104.9652
19.4577 | -104.9502
19.4613 | -104.9362
19.465 | -104.9224
19.4687 | -104.9079
19.4723 | -104.8936
19.476 | -104.8819
19.4796 | -104.8702
19.4833 | -104.8568
19.4869 | -104.8454
19.4906 | -104.8376
19.4943 | -104.8286
19.4979 | -104.8186
19.5016 | -104.8083
19.5052 | -104.8006
19.5089 | -104.7957
19.5126 | -104.7923
19.5162 | -104.79
19.5199 | -104.79
19.5235 | -104.8027
19.5272 | -104.8039
19.5308 | -104.8038
19.5345 | -104.8082
19.5382 | -104.814
19.5418 | -104.8202
19.5455 | -104.8288
19.5491 | -104.8388
19.5528 | -104.8505
19.5565 | -104.8645
19.5601 | -104.88
19.5638 | -104.8966
19.5674 | -104.9109
19.5711 | -104.9263
19.5747 | -104.9409
19.5784 | -104.9594
19.5821 | -104.979
19.5857 | -104.9968
19.5894 | -105.0167
19.593 | -105.037
19.5967 | -105.0589
19.6003 | -105.0799
19.604 | -105.1013
19.6077 | -105.1209
19.6113 | -105.1393
19.615 | -105.16
19.6186 | -105.1779
19.6223 | -105.1921
19.626 | -105.2064
19.6296 | -105.22
19.6333 | -105.2369
19.6369 | -105.2534
19.6406 | -105.268
19.6442 | -105.2806
19.6479 | -105.2907
19.6516 | -105.2991
19.6552 | -105.3067
19.6589 | -105.313
19.6625 | -105.3178
19.6662 | -105.321
19.6699 | -105.3224
19.6735 | -105.3242
19.6772 | -105.3252
19.6808 | -105.323
19.6845 | -105.32
19.6881 | -105.3166
19.6918 | -105.311
19.6955 | -105.3033
19.6991 | -105.2943
19.7028 | -105.2821
19.7064 | -105.27
19.7101 | -105.258
19.7138 | -105.2457
19.7174 | -105.2324
19.7211 | -105.2191
19.7247 | -105.2055
19.7284 | -105.1929
19.732 | -105.1789
19.7357 | -105.1631
19.7394 | -105.1486
19.743 | -105.1364
19.7467 | -105.1254
19.7503 | -105.1128
19.754 | -105.1012
19.7576 | -105.0918
19.7613 | -105.0826
19.765 | -105.072
19.7686 | -105.0603
19.7723 | -105.0493
19.7759 | -105.0402
19.7796 | -105.0342
19.7833 | -105.028
19.7869 | -105.0191
19.7906 | -105.0124
19.7942 | -105.0073
19.7979 | -105.0029
19.8015 | -104.9968
19.8052 | -104.9918
19.8089 | -104.9912
19.8125 | -104.9922
19.8162 | -104.9917
19.8198 | -104.9932
19.8235 | -104.9936
19.8272 | -104.9923
19.8308 | -104.9918
19.8345 | -104.9917
19.8381 | -104.9957
19.8418 | -105.0013
19.8454 | -105.0047
19.8491 | -105.0098
19.8528 | -105.0152
19.8564 | -105.021
19.8601 | -105.0274
19.8637 | -105.0343
19.8674 | -105.0417
19.8711 | -105.0481
19.8747 | -105.0547
19.8784 | -105.0601
19.882 | -105.0645
19.8857 | -105.0705
19.8893 | -105.0785
19.893 | -105.0842
19.8967 | -105.0862
19.9003 | -105.0875
19.904 | -105.0919
19.9076 | -105.0932
19.9113 | -105.0943
19.9149 | -105.0929
19.9186 | -105.0912
19.9223 | -105.0892
19.9259 | -105.0859
19.9296 | -105.082
19.9332 | -105.0773
19.9369 | -105.0719
19.9406 | -105.0662
19.9442 | -105.0584
19.9479 | -105.0499
19.9515 | -105.0419
19.9552 | -105.0333
19.9588 | -105.0246
19.9625 | -105.0168
19.9662 | -105.0071
19.9698 | -104.9944
19.9735 | -104.9837
19.9771 | -104.9757
19.9808 | -104.9673
19.9845 | -104.9585
19.9881 | -104.9496
19.9918 | -104.9414
19.9954 | -104.9324
19.9991 | -104.9229
20.0027 | -104.9142
20.0064 | -104.9065
20.0101 | -104.8996
20.0137 | -104.8957
20.0174 | -104.8927
20.021 | -104.8899
20.0247 | -104.8838
20.0284 | -104.8794
20.032 | -104.8777
20.0357 | -104.878
20.0393 | -104.8795
20.043 | -104.8815
20.0466 | -104.8854
20.0503 | -104.8906
20.054 | -104.8967
20.0576 | -104.8988
20.0613 | -104.9021
20.0649 | -104.9053
20.0686 | -104.9072
20.0722 | -104.9109
20.0759 | -104.9149
20.0796 | -104.9187
20.0832 | -104.9211
20.0869 | -104.9261
20.0905 | -104.9323
20.0942 | -104.9363
20.0979 | -104.94
20.1015 | -104.943
20.1052 | -104.9447
20.1088 | -104.9467
20.1125 | -104.9501
20.1161 | -104.9534
20.1198 | -104.9582
20.1235 | -104.9587
20.1271 | -104.9578
20.1308 | -104.9595
20.1344 | -104.9622
20.1381 | -104.9642
20.1418 | -104.9659
20.1454 | -104.966
20.1491 | -104.9667
20.1527 | -104.967
20.1564 | -104.9665
20.16 | -104.9685
20.1637 | -104.9712
20.1674 | -104.9745
20.171 | -104.9753
20.1747 | -104.9744
20.1783 | -104.9751
20.182 | -104.9769
20.1857 | -104.9807
20.1893 | -104.9822
20.193 | -104.9851
20.1966 | -104.9877
20.2003 | -104.9882
20.2039 | -104.9866
20.2076 | -104.9858
20.2113 | -104.9853
20.2149 | -104.9851
20.2186 | -104.9864
20.2222 | -104.9858
20.2259 | -104.9861
20.2295 | -104.9888
20.2332 | -104.9911
20.2369 | -104.9896
20.2405 | -104.9874
20.2442 | -104.9868
20.2478 | -104.9856
20.2515 | -104.985
20.2552 | -104.9846
20.2588 | -104.982
20.2625 | -104.9781
20.2661 | -104.9744
20.2698 | -104.9721
20.2734 | -104.9708
20.2771 | -104.9682
20.2808 | -104.9648
20.2844 | -104.9603
20.2881 | -104.9526
20.2917 | -104.9448
20.2954 | -104.9373
20.2991 | -104.9301
20.3027 | -104.9239
20.3064 | -104.9172
20.31 | -104.9105
20.3137 | -104.9008
20.3173 | -104.8919
20.321 | -104.8809
20.3247 | -104.8701
20.3283 | -104.8593
20.332 | -104.8468
20.3356 | -104.8344
20.3393 | -104.8244
20.343 | -104.8157
20.3466 | -104.806
20.3503 | -104.792
20.3539 | -104.779
20.3576 | -104.7672
20.3612 | -104.7538
20.3649 | -104.7392
20.3686 | -104.7259
20.3722 | -104.7143
20.3759 | -104.7011
20.3795 | -104.6861
20.3832 | -104.6728
20.3868 | -104.6608
20.3905 | -104.6493
20.3942 | -104.6385
20.3978 | -104.6301
20.4015 | -104.6225
20.4051 | -104.6145
20.4088 | -104.6065
20.4125 | -104.6018
20.4161 | -104.5982
20.4198 | -104.5946
20.4234 | -104.5915
20.4271 | -104.5896
20.4307 | -104.5882
20.4344 | -104.5893
20.4381 | -104.5909
20.4417 | -104.5922
20.4454 | -104.5938
20.449 | -104.5934
20.4527 | -104.5951
20.4564 | -104.5986
20.46 | -104.6022
20.4637 | -104.6059
20.4673 | -104.6108
20.471 | -104.6151
20.4746 | -104.6196
20.4783 | -104.6243
20.482 | -104.6283
20.4856 | -104.6337
20.4893 | -104.6375
20.4929 | -104.6415
20.4966 | -104.6458
20.5003 | -104.648
20.5039 | -104.6504
20.5076 | -104.6522
20.5112 | -104.6533
20.5149 | -104.6553
20.5185 | -104.6564
20.5222 | -104.6552
20.5259 | -104.6571
20.5295 | -104.659
20.5332 | -104.658
20.5368 | -104.6551
20.5405 | -104.652
20.5441 | -104.6503
20.5478 | -104.6473
20.5515 | -104.6435
20.5551 | -104.6381
20.5588 | -104.6326
20.5624 | -104.6272
20.5661 | -104.621
20.5698 | -104.6158
20.5734 | -104.61
20.5771 | -104.6078
20.5807 | -104.6008
20.5844 | -104.5965
20.588 | -104.5916
20.5917 | -104.5854
20.5954 | -104.5773
20.599 | -104.5686
20.6027 | -104.5619
20.6063 | -104.5534
20.61 | -104.5453
20.6137 | -104.5386
20.6173 | -104.5318
20.621 | -104.5265
20.6246 | -104.5212
20.6283 | -104.5138
20.6319 | -104.5057
20.6356 | -104.4993
20.6393 | -104.4947
20.6429 | -104.4905
20.6466 | -104.4859
20.6502 | -104.4821
20.6539 | -104.4797
20.6576 | -104.4766
20.6612 | -104.4732
20.6649 | -104.4701
20.6685 | -104.4672
20.6722 | -104.4656
20.6758 | -104.4643
20.6795 | -104.4628
20.6832 | -104.4602
20.6868 | -104.4584
20.6905 | -104.4567
20.6941 | -104.4545
20.6978 | -104.4522
20.7015 | -104.4497
20.7051 | -104.447
20.7088 | -104.4472
20.7124 | -104.4468
20.7161 | -104.4474
20.7197 | -104.4482
20.7234 | -104.4489
20.7271 | -104.4498
20.7307 | -104.4501
20.7344 | -104.4491
20.738 | -104.4501
20.7417 | -104.4529
20.7453 | -104.4557
20.749 | -104.4567
20.7527 | -104.4581
20.7563 | -104.4609
20.76 | -104.463
20.7636 | -104.4657
20.7673 | -104.4683
20.771 | -104.4709
20.7746 | -104.4738
20.7783 | -104.4733
20.7819 | -104.4764
20.7856 | -104.4789
20.7892 | -104.4801
20.7929 | -104.4832
20.7966 | -104.4855
20.8002 | -104.4876
20.8039 | -104.4894
20.8075 | -104.4877
20.8112 | -104.4872
20.8149 | -104.4892
20.8185 | -104.4905
20.8222 | -104.4904
20.8258 | -104.4895
20.8295 | -104.4909
20.8331 | -104.4932
20.8368 | -104.4942
20.8405 | -104.4961
20.8441 | -104.4956
20.8478 | -104.4962
20.8514 | -104.4978
20.8551 | -104.4978
20.8588 | -104.4996
20.8624 | -104.5011
20.8661 | -104.5041
20.8697 | -104.5053
20.8734 | -104.5065
20.877 | -104.5106
20.8807 | -104.5149
20.8844 | -104.518
20.888 | -104.5212
20.8917 | -104.5255
20.8953 | -104.5318
20.899 | -104.5385
20.9026 | -104.5469
20.9063 | -104.5533
20.91 | -104.5571
20.9136 | -104.5633
20.9173 | -104.5722
20.9209 | -104.5803
20.9246 | -104.5899
20.9283 | -104.5989
20.9319 | -104.6046
20.9356 | -104.6107
20.9392 | -104.6191
20.9429 | -104.6263
20.9465 | -104.6318
20.9502 | -104.6383
20.9539 | -104.6461
20.9575 | -104.652
20.9612 | -104.6568
20.9648 | -104.6623
20.9685 | -104.6658
20.9722 | -104.669
20.9758 | -104.6708
20.9795 | -104.6726
20.9831 | -104.6711
20.9868 | -104.6665
20.9904 | -104.6601
20.9941 | -104.6568
20.9978 | -104.6527
21.0014 | -104.6462
21.0051 | -104.6363
21.0087 | -104.6265
21.0124 | -104.6133
21.016 | -104.5981
21.0197 | -104.5834
21.0234 | -104.5665
21.027 | -104.5487
21.0307 | -104.5331
21.0343 | -104.5156
21.038 | -104.4966
21.0417 | -104.4746
21.0453 | -104.454
21.049 | -104.4317
21.0526 | -104.4084
21.0563 | -104.3853
21.0599 | -104.3605
21.0636 | -104.3343
21.0673 | -104.3078
21.0709 | -104.2813
21.0746 | -104.2553
21.0782 | -104.2284
21.0819 | -104.2024
21.0856 | -104.1755
21.0892 | -104.1475
21.0929 | -104.118
21.0965 | -104.0876
21.1002 | -104.0609
21.1038 | -104.0347
21.1075 | -104.004
21.1112 | -103.9761
21.1148 | -103.9476
21.1185 | -103.9184
21.1221 | -103.8903
21.1258 | -103.8638
21.1295 | -103.8371
21.1331 | -103.8101
21.1368 | -103.784
21.1404 | -103.7583
21.1441 | -103.7353
21.1477 | -103.712
21.1514 | -103.6896
21.1551 | -103.6687
21.1587 | -103.6485
21.1624 | -103.6293
21.166 | -103.6115
21.1697 | -103.5945
21.1733 | -103.58
21.177 | -103.5662
21.1807 | -103.5534
21.1843 | -103.5434
21.188 | -103.5357
21.1916 | -103.5298
21.1953 | -103.5242
21.199 | -103.5195
21.2026 | -103.5179
21.2063 | -103.5172
21.2099 | -103.5182
21.2136 | -103.5196
21.2172 | -103.5208
21.2209 | -103.5239
21.2246 | -103.528
21.2282 | -103.5335
21.2319 | -103.5411
21.2355 | -103.5506
21.2392 | -103.5606
21.2429 | -103.5713
21.2465 | -103.5822
21.2502 | -103.5927
21.2538 | -103.6052
21.2575 | -103.6193
21.2611 | -103.6347
21.2648 | -103.6501
21.2685 | -103.6665
21.2721 | -103.6824
21.2758 | -103.7003
21.2794 | -103.7182
21.2831 | -103.7371
21.2868 | -103.7569
21.2904 | -103.7762
21.2941 | -103.7945
21.2977 | -103.8132
21.3014 | -103.8322
21.305 | -103.8506
21.3087 | -103.8674
21.3124 | -103.8841
21.316 | -103.9026
21.3197 | -103.9194
21.3233 | -103.9357
21.327 | -103.9517
21.3307 | -103.9653
21.3343 | -103.9797
21.338 | -103.9928
21.3416 | -104.0054
21.3453 | -104.0155
21.3489 | -104.0258
21.3526 | -104.0351
21.3563 | -104.0433
21.3599 | -104.0516
21.3636 | -104.0595
21.3672 | -104.0666
21.3709 | -104.0734
21.3745 | -104.0799
21.3782 | -104.0865
21.3819 | -104.0915
21.3855 | -104.0949
21.3892 | -104.0977
21.3928 | -104.1012
21.3965 | -104.1043
21.4002 | -104.1048
21.4038 | -104.1057
21.4075 | -104.1058
21.4111 | -104.1062
21.4148 | -104.107
21.4184 | -104.1074
21.4221 | -104.1067
21.4258 | -104.1058
21.4294 | -104.1035
21.4331 | -104.1008
21.4367 | -104.0981
21.4404 | -104.095
21.4441 | -104.0939
21.4477 | -104.0929
21.4514 | -104.0896
21.455 | -104.0826
21.4587 | -104.0772
21.4623 | -104.0732
21.466 | -104.0705
21.4697 | -104.0678
21.4733 | -104.0661
21.477 | -104.0614
21.4806 | -104.0572
21.4843 | -104.0497
21.488 | -104.0419
21.4916 | -104.037
21.4953 | -104.0324
21.4989 | -104.0267
21.5026 | -104.0203
21.5062 | -104.0131
21.5099 | -104.006
21.5136 | -103.9999
21.5172 | -103.9935
21.5209 | -103.9873
21.5245 | -103.98
21.5282 | -103.9718
21.5318 | -103.9637
21.5355 | -103.9544
21.5392 | -103.9428
21.5428 | -103.9293
21.5465 | -103.9145
21.5501 | -103.9011
21.5538 | -103.888
21.5575 | -103.8723
21.5611 | -103.8565
21.5648 | -103.84
21.5684 | -103.8231
21.5721 | -103.8066
21.5757 | -103.79
21.5794 | -103.7722
21.5831 | -103.7551
21.5867 | -103.7384
21.5904 | -103.7214
21.594 | -103.7024
21.5977 | -103.6811
21.6014 | -103.6609
21.605 | -103.6413
21.6087 | -103.623
21.6123 | -103.6055
21.616 | -103.5888
21.6196 | -103.5715
21.6233 | -103.5547
21.627 | -103.5398
21.6306 | -103.5245
21.6343 | -103.5093
21.6379 | -103.4987
21.6416 | -103.4841
21.6453 | -103.4709
21.6489 | -103.4602
21.6526 | -103.4515
21.6562 | -103.4433
21.6599 | -103.4351
21.6635 | -103.4278
21.6672 | -103.4213
21.6709 | -103.4166
21.6745 | -103.4125
21.6782 | -103.4089
21.6818 | -103.4065
21.6855 | -103.4045
21.6891 | -103.4042
21.6928 | -103.4065
21.6965 | -103.4106
21.7001 | -103.4149
21.7038 | -103.4204
21.7074 | -103.4284
21.7111 | -103.4375
21.7148 | -103.4484
21.7184 | -103.4603
21.7221 | -103.472
21.7257 | -103.4862
21.7294 | -103.5031
21.733 | -103.5205
21.7367 | -103.5402
21.7404 | -103.5609
21.744 | -103.5839
21.7477 | -103.6074
21.7513 | -103.6299
21.755 | -103.6516
21.7587 | -103.6769
21.7623 | -103.7044
21.766 | -103.7299
21.7696 | -103.7542
21.7733 | -103.7798
21.7769 | -103.8072
21.7806 | -103.8335
21.7843 | -103.8615
21.7879 | -103.8893
21.7916 | -103.9196
21.7952 | -103.9478
21.7989 | -103.9723
21.8026 | -103.9957
21.8062 | -104.0202
21.8099 | -104.0435
21.8135 | -104.0634
21.8172 | -104.0824
21.8208 | -104.1015
21.8245 | -104.1185
21.8282 | -104.1337
21.8318 | -104.1511
21.8355 | -104.1674
21.8391 | -104.1798
21.8428 | -104.193
21.8464 | -104.2052
21.8501 | -104.2158
21.8538 | -104.2261
21.8574 | -104.2331
21.8611 | -104.2364
21.8647 | -104.2383
21.8684 | -104.2411
21.8721 | -104.2431
21.8757 | -104.2427
21.8794 | -104.2424
21.883 | -104.2423
21.8867 | -104.2404
21.8903 | -104.2365
21.894 | -104.2332
21.8977 | -104.2303
21.9013 | -104.2301
21.905 | -104.2282
21.9086 | -104.2245
21.9123 | -104.2204
21.916 | -104.2158
21.9196 | -104.211
21.9233 | -104.2063
21.9269 | -104.2001
21.9306 | -104.1945
21.9342 | -104.1875
21.9379 | -104.1801
21.9416 | -104.1724
21.9452 | -104.1674
21.9489 | -104.1612
21.9525 | -104.1561
21.9562 | -104.1501
21.9599 | -104.143
21.9635 | -104.1372
21.9672 | -104.1298
21.9708 | -104.124
21.9745 | -104.1205
21.9781 | -104.1185
21.9818 | -104.1154
21.9855 | -104.1125
21.9891 | -104.11
21.9928 | -104.1081
21.9964 | -104.1071
22.0001 | -104.1056
22.0038 | -104.1044
22.0074 | -104.1026
22.0111 | -104.1023
22.0147 | -104.1018
22.0184 | -104.1004
22.022 | -104.1
22.0257 | -104.0991
22.0294 | -104.0985
22.033 | -104.0981
22.0367 | -104.0958
22.0403 | -104.0921
22.044 | -104.0883
22.0476 | -104.0851
22.0513 | -104.0826
22.055 | -104.0781
22.0586 | -104.0739
22.0623 | -104.069
22.0659 | -104.0641
22.0696 | -104.0591
22.0733 | -104.0551
22.0769 | -104.0503
22.0806 | -104.0444
22.0842 | -104.0347
22.0879 | -104.0241
22.0915 | -104.0152
22.0952 | -104.0058
22.0989 | -103.9952
22.1025 | -103.9822
22.1062 | -103.9707
22.1098 | -103.961
22.1135 | -103.9501
22.1172 | -103.9386
22.1208 | -103.9295
22.1245 | -103.9202
22.1281 | -103.9099
22.1318 | -103.8979
22.1354 | -103.8874
22.1391 | -103.8768
22.1428 | -103.8639
22.1464 | -103.8523
22.1501 | -103.8429
22.1537 | -103.8331
22.1574 | -103.8216
22.161 | -103.8127
22.1647 | -103.8081
22.1684 | -103.8029
22.172 | -103.7974
22.1757 | -103.7937
22.1793 | -103.7902
22.183 | -103.7844
22.1867 | -103.779
22.1903 | -103.7751
22.194 | -103.7709
22.1976 | -103.7674
22.2013 | -103.7649
22.2049 | -103.763
22.2086 | -103.7611
22.2123 | -103.7607
22.2159 | -103.7609
22.2196 | -103.7632
22.2232 | -103.7652
22.2269 | -103.7666
22.2306 | -103.7697
22.2342 | -103.7736
22.2379 | -103.7775
22.2415 | -103.7838
22.2452 | -103.7885
22.2488 | -103.7945
22.2525 | -103.8018
22.2562 | -103.8105
22.2598 | -103.8155
22.2635 | -103.8203
22.2671 | -103.8263
22.2708 | -103.8335
22.2745 | -103.8393
22.2781 | -103.8441
22.2818 | -103.8514
22.2854 | -103.8577
22.2891 | -103.8616
22.2927 | -103.8654
22.2964 | -103.87
22.3001 | -103.8738
22.3037 | -103.8785
22.3074 | -103.8841
22.311 | -103.8891
22.3147 | -103.8948
22.3183 | -103.8991
22.322 | -103.9036
22.3257 | -103.9066
22.3293 | -103.9105
22.333 | -103.9142
22.3366 | -103.9173
22.3403 | -103.9217
22.344 | -103.9256
22.3476 | -103.9298
22.3513 | -103.9362
22.3549 | -103.9426
22.3586 | -103.9488
22.3622 | -103.9542
22.3659 | -103.9592
22.3696 | -103.9649
22.3732 | -103.9701
22.3769 | -103.9761
22.3805 | -103.9848
22.3842 | -103.992
22.3879 | -103.9977
22.3915 | -104.0056
22.3952 | -104.0143
22.3988 | -104.0228
22.4025 | -104.0319
22.4061 | -104.0417
22.4098 | -104.0518
22.4135 | -104.0616
22.4171 | -104.0742
22.4208 | -104.0868
22.4244 | -104.0961
22.4281 | -104.1076
22.4318 | -104.1198
22.4354 | -104.1305
22.4391 | -104.1418
22.4427 | -104.1444
22.4464 | -104.1348
22.45 | -104.1174
22.4537 | -104.1051
22.4574 | -104.0924
22.461 | -104.0694
22.4647 | -104.0482
22.4683 | -104.0325
22.472 | -104.019
22.4757 | -104.0028
22.4793 | -103.9838
22.483 | -103.9604
22.4866 | -103.9337
22.4903 | -103.9069
22.4939 | -103.8774
22.4976 | -103.8438
22.5013 | -103.8057
22.5049 | -103.7654
22.5086 | -103.7235
22.5122 | -103.6797
22.5159 | -103.6339
22.5195 | -103.5862
22.5232 | -103.5367
22.5269 | -103.485
22.5305 | -103.4314
22.5342 | -103.377
22.5378 | -103.3219
22.5415 | -103.2668
22.5452 | -103.2119
22.5488 | -103.1558
22.5525 | -103.1014
22.5561 | -103.0491
22.5598 | -102.997
22.5634 | -102.9465
22.5671 | -102.8971
22.5708 | -102.8504
22.5744 | -102.8077
22.5781 | -102.7695
22.5817 | -102.7331
22.5854 | -102.6962
22.5891 | -102.664
22.5927 | -102.6373
22.5964 | -102.6132
22.6 | -102.5925
22.6037 | -102.5753
22.6073 | -102.5634
22.611 | -102.5549
22.6147 | -102.5517
22.6183 | -102.5537
22.622 | -102.5577
22.6256 | -102.5656
22.6293 | -102.5771
22.633 | -102.592
22.6366 | -102.6109
22.6403 | -102.6337
22.6439 | -102.6597
22.6476 | -102.6877
22.6512 | -102.7198
22.6549 | -102.7532
22.6586 | -102.7888
22.6622 | -102.8273
22.6659 | -102.8673
22.6695 | -102.9084
22.6732 | -102.9522
22.6768 | -102.9981
22.6805 | -103.0438
22.6842 | -103.0891
22.6878 | -103.1347
22.6915 | -103.1726
22.6951 | -103.2156
22.6988 | -103.2585
22.7025 | -103.2996
22.7061 | -103.3399
22.7098 | -103.3792
22.7134 | -103.4159
22.7171 | -103.4517
22.7207 | -103.4889
22.7244 | -103.5199
22.7281 | -103.5479
22.7317 | -103.5762
22.7354 | -103.6033
22.739 | -103.6289
22.7427 | -103.6524
22.7464 | -103.6742
22.75 | -103.6923
22.7537 | -103.7081
22.7573 | -103.7227
22.761 | -103.7364
22.7646 | -103.7575
22.7683 | -103.7849
22.772 | -103.8118
22.7756 | -103.8276
22.7793 | -103.8382
22.7829 | -103.8542
22.7866 | -103.8627
22.7903 | -103.8617
22.7939 | -103.8576
22.7976 | -103.8513
22.8012 | -103.8467
22.8049 | -103.8445
22.8085 | -103.8405
22.8122 | -103.8351
22.8159 | -103.8286
22.8195 | -103.8246
22.8232 | -103.8229
22.8268 | -103.8204
22.8305 | -103.8172
22.8341 | -103.8136
22.8378 | -103.8076
22.8415 | -103.8006
22.8451 | -103.7944
22.8488 | -103.7898
22.8524 | -103.7861
22.8561 | -103.7815
22.8598 | -103.7751
22.8634 | -103.7678
22.8671 | -103.7597
22.8707 | -103.7517
22.8744 | -103.7439
22.878 | -103.7364
22.8817 | -103.729
22.8854 | -103.7216
22.889 | -103.7138
22.8927 | -103.705
22.8963 | -103.6934
22.9 | -103.6808
22.9037 | -103.6684
22.9073 | -103.6577
22.911 | -103.6443
22.9146 | -103.6309
22.9183 | -103.6183
22.9219 | -103.6036
22.9256 | -103.5878
22.9293 | -103.5713
22.9329 | -103.5562
22.9366 | -103.536
22.9402 | -103.516
22.9439 | -103.4988
22.9476 | -103.4799
22.9512 | -103.4626
22.9549 | -103.4454
22.9585 | -103.4283
22.9622 | -103.4076
22.9658 | -103.3854
22.9695 | -103.3647
22.9732 | -103.3424
22.9768 | -103.3224
22.9805 | -103.3018
22.9841 | -103.2803
22.9878 | -103.26
22.9914 | -103.2382
22.9951 | -103.2153
22.9988 | -103.1917
23.0024 | -103.1704
23.0061 | -103.15
23.0097 | -103.1308
23.0134 | -103.1108
23.0171 | -103.0904
23.0207 | -103.0715
23.0244 | -103.0515
23.028 | -103.0312
23.0317 | -103.0116
23.0353 | -102.9898
23.039 | -102.9694
23.0427 | -102.9519
23.0463 | -102.9321
23.05 | -102.9136
23.0536 | -102.8961
23.0573 | -102.8764
23.061 | -102.8548
23.0646 | -102.8347
23.0683 | -102.8145
23.0719 | -102.7968
23.0756 | -102.7805
23.0792 | -102.7632
23.0829 | -102.7465
23.0866 | -102.7303
23.0902 | -102.7128
23.0939 | -102.6943
23.0975 | -102.6774
23.1012 | -102.6603
23.1049 | -102.6439
23.1085 | -102.6288
23.1122 | -102.6157
23.1158 | -102.603
23.1195 | -102.5904
23.1231 | -102.5765
23.1268 | -102.5608
23.1305 | -102.5462
23.1341 | -102.5328
23.1378 | -102.5186
23.1414 | -102.5056
23.1451 | -102.4932
23.1487 | -102.4826
23.1524 | -102.4721
23.1561 | -102.4623
23.1597 | -102.4507
23.1634 | -102.4399
23.167 | -102.4303
23.1707 | -102.4207
23.1744 | -102.4103
23.178 | -102.3995
23.1817 | -102.389
23.1853 | -102.378
23.189 | -102.3688
23.1926 | -102.3607
23.1963 | -102.3533
23.2 | -102.3442
23.2036 | -102.3334
23.2073 | -102.3229
23.2109 | -102.3137
23.2146 | -102.3035
23.2183 | -102.2926
23.2219 | -102.2818
23.2256 | -102.2723
23.2292 | -102.2627
23.2329 | -102.2533
23.2365 | -102.2448
23.2402 | -102.2384
23.2439 | -102.2304
23.2475 | -102.2217
23.2512 | -102.2137
23.2548 | -102.207
23.2585 | -102.2002
23.2622 | -102.1918
23.2658 | -102.1855
23.2695 | -102.1795
23.2731 | -102.1732
23.2768 | -102.1669
23.2804 | -102.159
23.2841 | -102.1512
23.2878 | -102.1441
23.2914 | -102.1406
23.2951 | -102.1391
23.2987 | -102.137
23.3024 | -102.1334
23.306 | -102.1288
23.3097 | -102.1249
23.3134 | -102.1213
23.317 | -102.1191
23.3207 | -102.1165
23.3243 | -102.1147
23.328 | -102.1132
23.3317 | -102.1124
23.3353 | -102.1122
23.339 | -102.1094
23.3426 | -102.1086
23.3463 | -102.1087
23.3499 | -102.1103
23.3536 | -102.1127
23.3573 | -102.1148
23.3609 | -102.1184
23.3646 | -102.1235
23.3682 | -102.1281
23.3719 | -102.131
23.3756 | -102.1355
23.3792 | -102.1418
23.3829 | -102.147
23.3865 | -102.1532
23.3902 | -102.1613
23.3938 | -102.1671
23.3975 | -102.1731
23.4012 | -102.1799
23.4048 | -102.1881
23.4085 | -102.1959
23.4121 | -102.2024
23.4158 | -102.2101
23.4195 | -102.2189
23.4231 | -102.2285
23.4268 | -102.237
23.4304 | -102.248
23.4341 | -102.2587
23.4377 | -102.2673
23.4414 | -102.2756
23.4451 | -102.284
23.4487 | -102.2927
23.4524 | -102.304
23.456 | -102.3169
23.4597 | -102.3304
23.4633 | -102.3434
23.467 | -102.3563
23.4707 | -102.3691
23.4743 | -102.381
23.478 | -102.3927
23.4816 | -102.4049
23.4853 | -102.4166
23.489 | -102.4295
23.4926 | -102.4437
23.4963 | -102.4574
23.4999 | -102.4711
23.5036 | -102.4844
23.5072 | -102.4974
23.5109 | -102.5106
23.5146 | -102.5255
23.5182 | -102.5391
23.5219 | -102.5515
23.5255 | -102.5621
23.5292 | -102.574
23.5329 | -102.5871
23.5365 | -102.6006
23.5402 | -102.6134
23.5438 | -102.6287
23.5475 | -102.643
23.5511 | -102.6573
23.5548 | -102.6703
23.5585 | -102.6821
23.5621 | -102.6939
23.5658 | -102.7069
23.5694 | -102.7192
23.5731 | -102.7275
23.5768 | -102.7367
23.5804 | -102.748
23.5841 | -102.7596
23.5877 | -102.7705
23.5914 | -102.7802
23.595 | -102.7894
23.5987 | -102.8004
23.6024 | -102.8107
23.606 | -102.8194
23.6097 | -102.8292
23.6133 | -102.84
23.617 | -102.8505
23.6207 | -102.8608
23.6243 | -102.8674
23.628 | -102.8746
23.6316 | -102.8815
23.6353 | -102.8894
23.6389 | -102.8983
23.6426 | -102.9052
23.6463 | -102.9114
23.6499 | -102.9151
23.6536 | -102.9204
23.6572 | -102.9251
23.6609 | -102.9307
23.6645 | -102.9345
23.6682 | -102.9372
23.6719 | -102.9426
23.6755 | -102.9469
23.6792 | -102.95
23.6828 | -102.95
23.6865 | -102.9493
23.6902 | -102.9487
23.6938 | -102.947
23.6975 | -102.9455
23.7011 | -102.9434
23.7048 | -102.9411
23.7084 | -102.9377
23.7121 | -102.9333
23.7158 | -102.9299
23.7194 | -102.9264
23.7231 | -102.9217
23.7267 | -102.9173
23.7304 | -102.9123
23.7341 | -102.9056
23.7377 | -102.8988
23.7414 | -102.8922
23.745 | -102.8873
23.7487 | -102.8786
23.7523 | -102.8718
23.756 | -102.8646
23.7597 | -102.8565
23.7633 | -102.8493
23.767 | -102.8421
23.7706 | -102.8355
23.7743 | -102.8292
23.778 | -102.821
23.7816 | -102.813
23.7853 | -102.806
23.7889 | -102.7975
23.7926 | -102.7885
23.7962 | -102.778
23.7999 | -102.7668
23.8036 | -102.7583
23.8072 | -102.7509
23.8109 | -102.7422
23.8145 | -102.7334
23.8182 | -102.7265
23.8218 | -102.7193
23.8255 | -102.7128
23.8292 | -102.7048
23.8328 | -102.6984
23.8365 | -102.6919
23.8401 | -102.6867
23.8438 | -102.6807
23.8475 | -102.6751
23.8511 | -102.6696
23.8548 | -102.6653
23.8584 | -102.6612
23.8621 | -102.6559
23.8657 | -102.6519
23.8694 | -102.6479
23.8731 | -102.6447
23.8767 | -102.6416
23.8804 | -102.6373
23.884 | -102.6328
23.8877 | -102.6281
23.8914 | -102.6247
23.895 | -102.6208
23.8987 | -102.6168
23.9023 | -102.6122
23.906 | -102.6065
23.9096 | -102.602
23.9133 | -102.5973
23.917 | -102.59
23.9206 | -102.5822
23.9243 | -102.5741
23.9279 | -102.5673
23.9316 | -102.5604
23.9353 | -102.551
23.9389 | -102.5412
23.9426 | -102.5322
23.9462 | -102.5207
23.9499 | -102.5082
23.9535 | -102.4964
23.9572 | -102.4829
23.9609 | -102.4701
23.9645 | -102.4553
23.9682 | -102.4404
23.9718 | -102.4246
23.9755 | -102.4088
23.9791 | -102.3938
23.9828 | -102.3774
23.9865 | -102.3601
23.9901 | -102.3421
23.9938 | -102.3252
23.9974 | -102.3064
24.0011 | -102.2873
24.0048 | -102.2679
24.0084 | -102.2465
24.0121 | -102.2244
24.0157 | -102.2045
24.0194 | -102.1853
24.023 | -102.1654
24.0267 | -102.1456
24.0304 | -102.1241
24.034 | -102.1024
24.0377 | -102.0812
24.0413 | -102.0591
24.045 | -102.0346
24.0487 | -102.0104
24.0523 | -101.9882
24.056 | -101.966
24.0596 | -101.9427
24.0633 | -101.9198
24.0669 | -101.8971
24.0706 | -101.8742
24.0743 | -101.8495
24.0779 | -101.826
24.0816 | -101.8032
24.0852 | -101.7801
24.0889 | -101.7571
24.0926 | -101.7352
24.0962 | -101.7145
24.0999 | -101.694
24.1035 | -101.6735
24.1072 | -101.6521
24.1108 | -101.6311
24.1145 | -101.6111
24.1182 | -101.5916
24.1218 | -101.5733
24.1255 | -101.5551
24.1291 | -101.5377
24.1328 | -101.5195
24.1364 | -101.5028
24.1401 | -101.4871
24.1438 | -101.4718
24.1474 | -101.4563
24.1511 | -101.4398
24.1547 | -101.4244
24.1584 | -101.4096
24.1621 | -101.3963
24.1657 | -101.3833
24.1694 | -101.3721
24.173 | -101.3597
24.1767 | -101.347
24.1803 | -101.3344
24.184 | -101.3245
24.1877 | -101.3163
24.1913 | -101.3097
24.195 | -101.3028
24.1986 | -101.2966
24.2023 | -101.292
24.206 | -101.2868
24.2096 | -101.2805
24.2133 | -101.2748
24.2169 | -101.2699
24.2206 | -101.2652
24.2242 | -101.26
24.2279 | -101.2564
24.2316 | -101.2518
24.2352 | -101.2475
24.2389 | -101.2442
24.2425 | -101.2402
24.2462 | -101.2353
24.2499 | -101.2315
24.2535 | -101.2288
24.2572 | -101.2274
24.2608 | -101.2249
24.2645 | -101.2199
24.2681 | -101.2121
24.2718 | -101.1984
24.2755 | -101.1521
24.2791 | -101.1155
24.2828 | -101.0807
24.2864 | -101.0483
24.2901 | -101.0168
24.2937 | -100.9874
24.2974 | -100.9585
24.3011 | -100.9277
24.3047 | -100.8957
24.3084 | -100.8619
24.312 | -100.8253
24.3157 | -100.7879
24.3194 | -100.7466
24.323 | -100.7031
24.3267 | -100.6593
24.3303 | -100.6138
24.334 | -100.5671
24.3376 | -100.5181
24.3413 | -100.4678
24.345 | -100.4168
24.3486 | -100.364
24.3523 | -100.3128
24.3559 | -100.2611
24.3596 | -100.2061
24.3633 | -100.1518
24.3669 | -100.0986
24.3706 | -100.0455
24.3742 | -99.9932
24.3779 | -99.9406
24.3815 | -99.8892
24.3852 | -99.8397
24.3889 | -99.794
24.3925 | -99.7492
24.3962 | -99.7042
24.3998 | -99.6604
24.4035 | -99.6187
24.4072 | -99.5805
24.4108 | -99.5446
24.4145 | -99.5122
24.4181 | -99.4833
24.4218 | -99.4573
24.4254 | -99.4338
24.4291 | -99.4139
24.4328 | -99.3969
24.4364 | -99.3828
24.4401 | -99.3718
24.4437 | -99.3652
24.4474 | -99.3622
24.451 | -99.3634
24.4547 | -99.3668
24.4584 | -99.3736
24.462 | -99.3852
24.4657 | -99.4003
24.4693 | -99.4189
24.473 | -99.4409
24.4767 | -99.4669
24.4803 | -99.4957
24.484 | -99.5283
24.4876 | -99.564
24.4913 | -99.602
24.4949 | -99.6437
24.4986 | -99.6872
24.5023 | -99.7309
24.5059 | -99.7769
24.5096 | -99.826
24.5132 | -99.8767
24.5169 | -99.9287
24.5206 | -99.9825
24.5242 | -100.0373
24.5279 | -100.0903
24.5315 | -100.1441
24.5352 | -100.1973
24.5388 | -100.2507
24.5425 | -100.3046
24.5462 | -100.3571
24.5498 | -100.4093
24.5535 | -100.4618
24.5571 | -100.513
24.5608 | -100.56
24.5645 | -100.6024
24.5681 | -100.6411
24.5718 | -100.6709
24.5754 | -100.6927
24.5791 | -100.6992
24.5827 | -100.704
24.5864 | -100.7116
24.5901 | -100.7198
24.5937 | -100.7262
24.5974 | -100.7286
24.601 | -100.7272
24.6047 | -100.7215
24.6084 | -100.7121
24.612 | -100.7022
24.6157 | -100.6886
24.6193 | -100.6746
24.623 | -100.6612
24.6266 | -100.6476
24.6303 | -100.6243
24.634 | -100.5911
24.6376 | -100.5512
24.6413 | -100.5048
24.6449 | -100.4526
24.6486 | -100.3985
24.6522 | -100.342
24.6559 | -100.2845
24.6596 | -100.2261
24.6632 | -100.1661
24.6669 | -100.1074
24.6705 | -100.0474
24.6742 | -99.986
24.6779 | -99.9247
24.6815 | -99.8637
24.6852 | -99.8031
24.6888 | -99.7435
24.6925 | -99.6852
24.6961 | -99.6299
24.6998 | -99.5756
24.7035 | -99.5232
24.7071 | -99.4739
24.7108 | -99.4266
24.7144 | -99.3809
24.7181 | -99.3382
24.7218 | -99.2986
24.7254 | -99.2633
24.7291 | -99.2309
24.7327 | -99.2014
24.7364 | -99.1751
24.74 | -99.1527
24.7437 | -99.1344
24.7474 | -99.1211
24.751 | -99.1124
24.7547 | -99.1066
24.7583 | -99.1047
24.762 | -99.1066
24.7657 | -99.1137
24.7693 | -99.1244
24.773 | -99.1396
24.7766 | -99.1581
24.7803 | -99.1816
24.7839 | -99.2084
24.7876 | -99.2384
24.7913 | -99.2722
24.7949 | -99.309
24.7986 | -99.348
24.8022 | -99.4073
24.8059 | -99.4507
24.8095 | -99.496
24.8132 | -99.5436
24.8169 | -99.5926
24.8205 | -99.642
24.8242 | -99.692
24.8278 | -99.7429
24.8315 | -99.7964
24.8352 | -99.8508
24.8388 | -99.9047
24.8425 | -99.9599
24.8461 | -100.0158
24.8498 | -100.0705
24.8534 | -100.1248
24.8571 | -100.1786
24.8608 | -100.2312
24.8644 | -100.2826
24.8681 | -100.3337
24.8717 | -100.383
24.8754 | -100.4297
24.8791 | -100.4759
24.8827 | -100.5204
24.8864 | -100.5641
24.89 | -100.6054
24.8937 | -100.6453
24.8973 | -100.6839
24.901 | -100.7177
24.9047 | -100.7491
24.9083 | -100.7779
24.912 | -100.8086
24.9156 | -100.8408
24.9193 | -100.8729
24.923 | -100.9076
24.9266 | -100.9429
24.9303 | -100.9846
24.9339 | -101.0165
24.9376 | -101.0367
24.9412 | -101.0487
24.9449 | -101.0561
24.9486 | -101.0625
24.9522 | -101.0653
24.9559 | -101.0669
24.9595 | -101.0704
24.9632 | -101.07
24.9668 | -101.0734
24.9705 | -101.075
24.9742 | -101.0728
24.9778 | -101.0706
24.9815 | -101.0678
24.9851 | -101.0647
24.9888 | -101.0614
24.9925 | -101.0585
24.9961 | -101.0546
24.9998 | -101.0509
25.0034 | -101.0451
25.0071 | -101.0386
25.0107 | -101.0337
25.0144 | -101.0285
25.0181 | -101.0214
25.0217 | -101.0146
25.0254 | -101.0076
25.029 | -101.0009
25.0327 | -100.9932
25.0364 | -100.9849
25.04 | -100.9782
25.0437 | -100.9716
25.0473 | -100.9638
25.051 | -100.9554
25.0546 | -100.9461
25.0583 | -100.9371
25.062 | -100.9279
25.0656 | -100.9199
25.0693 | -100.9126
25.0729 | -100.9055
25.0766 | -100.8976
25.0803 | -100.8884
25.0839 | -100.8807
25.0876 | -100.873
25.0912 | -100.8651
25.0949 | -100.858
25.0985 | -100.8498
25.1022 | -100.8416
25.1059 | -100.8329
25.1095 | -100.826
25.1132 | -100.8185
25.1168 | -100.8087
25.1205 | -100.7999
25.1241 | -100.792
25.1278 | -100.7854
25.1315 | -100.7779
25.1351 | -100.7709
25.1388 | -100.764
25.1424 | -100.7572
25.1461 | -100.7494
25.1498 | -100.7427
25.1534 | -100.7383
25.1571 | -100.7336
25.1607 | -100.7267
25.1644 | -100.7216
25.168 | -100.7169
25.1717 | -100.712
25.1754 | -100.7084
25.179 | -100.705
25.1827 | -100.7009
25.1863 | -100.6945
25.19 | -100.6866
25.1937 | -100.6808
25.1973 | -100.6776
25.201 | -100.6721
25.2046 | -100.667
25.2083 | -100.6635
25.2119 | -100.66
25.2156 | -100.6556
25.2193 | -100.6496
25.2229 | -100.6433
25.2266 | -100.6374
25.2302 | -100.6336
25.2339 | -100.6289
25.2376 | -100.6235
25.2412 | -100.6161
25.2449 | -100.6097
25.2485 | -100.6024
25.2522 | -100.595
25.2558 | -100.5897
25.2595 | -100.5851
25.2632 | -100.5811
25.2668 | -100.5751
25.2705 | -100.5677
25.2741 | -100.5613
25.2778 | -100.5541
25.2814 | -100.5466
25.2851 | -100.5395
25.2888 | -100.5331
25.2924 | -100.5275
25.2961 | -100.5206
25.2997 | -100.5116
25.3034 | -100.5031
25.3071 | -100.496
25.3107 | -100.4893
25.3144 | -100.4807
25.318 | -100.4722
25.3217 | -100.4626
25.3253 | -100.453
25.329 | -100.445
25.3327 | -100.4373
25.3363 | -100.4279
25.34 | -100.4188
25.3436 | -100.4097
25.3473 | -100.3993
25.351 | -100.3881
25.3546 | -100.3784
25.3583 | -100.3671
25.3619 | -100.3595
25.3656 | -100.3514
25.3692 | -100.3405
25.3729 | -100.3297
25.3766 | -100.3207
25.3802 | -100.3124
25.3839 | -100.3056
25.3875 | -100.298
25.3912 | -100.2884
25.3949 | -100.2785
25.3985 | -100.2696
25.4022 | -100.2602
25.4058 | -100.2506
25.4095 | -100.2409
25.4131 | -100.2341
25.4168 | -100.2262
25.4205 | -100.2176
25.4241 | -100.2084
25.4278 | -100.1993
25.4314 | -100.1914
25.4351 | -100.1841
25.4387 | -100.1758
25.4424 | -100.1669
25.4461 | -100.1593
25.4497 | -100.1512
25.4534 | -100.1424
25.457 | -100.1349
25.4607 | -100.1284
25.4644 | -100.1213
25.468 | -100.1134
25.4717 | -100.106
25.4753 | -100.0991
25.479 | -100.0929
25.4826 | -100.0856
25.4863 | -100.0779
25.49 | -100.0713
25.4936 | -100.0651
25.4973 | -100.0607
25.5009 | -100.0563
25.5046 | -100.0521
25.5083 | -100.0456
25.5119 | -100.0378
25.5156 | -100.0314
25.5192 | -100.0269
25.5229 | -100.0216
25.5265 | -100.0163
25.5302 | -100.0129
25.5339 | -100.0091
25.5375 | -100.0056
25.5412 | -100.0046
25.5448 | -100.0019
25.5485 | -100
25.5522 | -99.9972
25.5558 | -99.9933
25.5595 | -99.9919
25.5631 | -99.9904
25.5668 | -99.9894
25.5704 | -99.9887
25.5741 | -99.9888
25.5778 | -99.9898
25.5814 | -99.9908
25.5851 | -99.9925
25.5887 | -99.9938
25.5924 | -99.9952
25.596 | -99.9965
25.5997 | -99.9987
25.6034 | -99.9997
25.607 | -100.0008
25.6107 | -100.0042
25.6143 | -100.008
25.618 | -100.0098
25.6217 | -100.0104
25.6253 | -100.0113
25.629 | -100.013
25.6326 | -100.014
25.6363 | -100.017
25.6399 | -100.021
25.6436 | -100.0237
25.6473 | -100.0261
25.6509 | -100.0285
25.6546 | -100.0323
25.6582 | -100.0367
25.6619 | -100.0397
25.6656 | -100.0428
25.6692 | -100.0453
25.6729 | -100.0474
25.6765 | -100.0498
25.6802 | -100.0529
25.6838 | -100.0562
25.6875 | -100.0593
25.6912 | -100.0606
25.6948 | -100.0631
25.6985 | -100.0683
25.7021 | -100.0715
25.7058 | -100.0746
25.7095 | -100.08
25.7131 | -100.0853
25.7168 | -100.0917
25.7204 | -100.0976
25.7241 | -100.1022
25.7277 | -100.1079
25.7314 | -100.1158
25.7351 | -100.1237
25.7387 | -100.1311
25.7424 | -100.1381
25.746 | -100.1441
25.7497 | -100.1505
25.7533 | -100.1604
25.757 | -100.172
25.7607 | -100.1811
25.7643 | -100.1899
25.768 | -100.2003
25.7716 | -100.2117
25.7753 | -100.2222
25.779 | -100.234
25.7826 | -100.2472
25.7863 | -100.2605
25.7899 | -100.2736
25.7936 | -100.2856
25.7972 | -100.2993
25.8009 | -100.3143
25.8046 | -100.3305
25.8082 | -100.3459
25.8119 | -100.3631
25.8155 | -100.3807
25.8192 | -100.3972
25.8229 | -100.4148
25.8265 | -100.4339
25.8302 | -100.4527
25.8338 | -100.4721
25.8375 | -100.4925
25.8411 | -100.513
25.8448 | -100.5342
25.8485 | -100.5562
25.8521 | -100.5764
25.8558 | -100.597
25.8594 | -100.6173
25.8631 | -100.6363
25.8668 | -100.6553
25.8704 | -100.6757
25.8741 | -100.6992
25.8777 | -100.7226
25.8814 | -100.7447
25.885 | -100.7666
25.8887 | -100.7875
25.8924 | -100.8066
25.896 | -100.8251
25.8997 | -100.8444
25.9033 | -100.8641
25.907 | -100.8813
25.9107 | -100.8999
25.9143 | -100.9182
25.918 | -100.9358
25.9216 | -100.9524
25.9253 | -100.9696
25.9289 | -100.9882
25.9326 | -101.0058
25.9363 | -101.0217
25.9399 | -101.0352
25.9436 | -101.0495
25.9472 | -101.0639
25.9509 | -101.0756
25.9545 | -101.0866
25.9582 | -101.0988
25.9619 | -101.1094
25.9655 | -101.1204
25.9692 | -101.1319
25.9728 | -101.1428
25.9765 | -101.1508
25.9802 | -101.1577
25.9838 | -101.1668
25.9875 | -101.1756
25.9911 | -101.1835
25.9948 | -101.1908
25.9984 | -101.1975
26.0021 | -101.2027
26.0058 | -101.2074
26.0094 | -101.2133
26.0131 | -101.2189
26.0167 | -101.2242
26.0204 | -101.2305
26.0241 | -101.2363
26.0277 | -101.2417
26.0314 | -101.2463
26.035 | -101.2506
26.0387 | -101.2557
26.0423 | -101.2614
26.046 | -101.2668
26.0497 | -101.2713
26.0533 | -101.2748
26.057 | -101.2793
26.0606 | -101.2837
26.0643 | -101.2866
26.0679 | -101.292
26.0716 | -101.2981
26.0753 | -101.3027
26.0789 | -101.3082
26.0826 | -101.312
26.0862 | -101.319
26.0899 | -101.3262
26.0936 | -101.3316
26.0972 | -101.3376
26.1009 | -101.3445
26.1045 | -101.3516
26.1082 | -101.3598
26.1118 | -101.3697
26.1155 | -101.3806
26.1192 | -101.3895
26.1228 | -101.3969
26.1265 | -101.4052
26.1301 | -101.4154
26.1338 | -101.4253
26.1375 | -101.4338
26.1411 | -101.4412
26.1448 | -101.4504
26.1484 | -101.4598
26.1521 | -101.4682
26.1557 | -101.4756
26.1594 | -101.484
26.1631 | -101.4925
26.1667 | -101.5019
26.1704 | -101.5083
26.174 | -101.5142
26.1777 | -101.5199
26.1814 | -101.5267
26.185 | -101.5335
26.1887 | -101.5391
26.1923 | -101.5444
26.196 | -101.5501
26.1996 | -101.5558
26.2033 | -101.5609
26.207 | -101.565
26.2106 | -101.5672
26.2143 | -101.5704
26.2179 | -101.5762
26.2216 | -101.5813
26.2252 | -101.5871
26.2289 | -101.5906
26.2326 | -101.5934
26.2362 | -101.5953
26.2399 | -101.599
26.2435 | -101.6013
26.2472 | -101.6017
26.2509 | -101.6019
26.2545 | -101.6022
26.2582 | -101.6024
26.2618 | -101.602
26.2655 | -101.6012
26.2691 | -101.6023
26.2728 | -101.6034
26.2765 | -101.6038
26.2801 | -101.6053
26.2838 | -101.6059
26.2874 | -101.6058
26.2911 | -101.6061
26.2948 | -101.6039
26.2984 | -101.6001
26.3021 | -101.5969
26.3057 | -101.5947
26.3094 | -101.5921
26.313 | -101.5877
26.3167 | -101.5854
26.3204 | -101.5848
26.324 | -101.5833
26.3277 | -101.5821
26.3313 | -101.5825
26.335 | -101.5818
26.3387 | -101.581
26.3423 | -101.5812
26.346 | -101.5809
26.3496 | -101.5794
26.3533 | -101.5779
26.3569 | -101.5758
26.3606 | -101.5749
26.3643 | -101.5733
26.3679 | -101.5705
26.3716 | -101.5689
26.3752 | -101.5674
26.3789 | -101.5655
26.3826 | -101.5655
26.3862 | -101.5656
26.3899 | -101.5647
26.3935 | -101.5633
26.3972 | -101.5611
26.4008 | -101.5575
26.4045 | -101.5564
26.4082 | -101.5579
26.4118 | -101.5576
26.4155 | -101.5558
26.4191 | -101.5541
26.4228 | -101.5525
26.4264 | -101.5522
26.4301 | -101.551
26.4338 | -101.5489
26.4374 | -101.5474
26.4411 | -101.5452
26.4447 | -101.5434
26.4484 | -101.5418
26.4521 | -101.5387
26.4557 | -101.5355
26.4594 | -101.5318
26.463 | -101.5262
26.4667 | -101.5213
26.4703 | -101.5156
26.474 | -101.5097
26.4777 | -101.5054
26.4813 | -101.5032
26.485 | -101.4988
26.4886 | -101.4934
26.4923 | -101.489
26.496 | -101.485
26.4996 | -101.4815
26.5033 | -101.477
26.5069 | -101.473
26.5106 | -101.4697
26.5142 | -101.4666
26.5179 | -101.461
26.5216 | -101.4557
26.5252 | -101.4514
26.5289 | -101.4483
26.5325 | -101.4447
26.5362 | -101.4401
26.5399 | -101.4362
26.5435 | -101.4324
26.5472 | -101.4285
26.5508 | -101.4248
26.5545 | -101.4199
26.5581 | -101.4164
26.5618 | -101.4132
26.5655 | -101.4093
26.5691 | -101.4062
26.5728 | -101.4028
26.5764 | -101.3988
26.5801 | -101.3967
26.5837 | -101.3958
26.5874 | -101.3937
26.5911 | -101.3924
26.5947 | -101.3906
26.5984 | -101.3892
26.602 | -101.3894
26.6057 | -101.3895
26.6094 | -101.3903
26.613 | -101.3895
26.6167 | -101.3896
26.6203 | -101.3927
26.624 | -101.3955
26.6276 | -101.397
26.6313 | -101.4002
26.635 | -101.4046
26.6386 | -101.4096
26.6423 | -101.4129
26.6459 | -101.4168
26.6496 | -101.422
26.6533 | -101.4278
26.6569 | -101.4348
26.6606 | -101.4424
26.6642 | -101.45
26.6679 | -101.4591
26.6715 | -101.4693
26.6752 | -101.4791
26.6789 | -101.489
26.6825 | -101.5015
26.6862 | -101.5148
26.6898 | -101.5257
26.6935 | -101.5361
26.6972 | -101.5482
26.7008 | -101.5592
26.7045 | -101.5696
26.7081 | -101.5802
26.7118 | -101.5905
26.7154 | -101.6003
26.7191 | -101.6109
26.7228 | -101.6231
26.7264 | -101.6338
26.7301 | -101.6467
26.7337 | -101.6589
26.7374 | -101.6699
26.741 | -101.6808
26.7447 | -101.6909
26.7484 | -101.6988
26.752 | -101.7043
26.7557 | -101.7109
26.7593 | -101.7183
26.763 | -101.726
26.7667 | -101.7339
26.7703 | -101.7395
26.774 | -101.7452
26.7776 | -101.7519
26.7813 | -101.7561
26.7849 | -101.7595
26.7886 | -101.7635
26.7923 | -101.7673
26.7959 | -101.7704
26.7996 | -101.7728
26.8032 | -101.7757
26.8069 | -101.7794
26.8106 | -101.783
26.8142 | -101.7848
26.8179 | -101.7862
26.8215 | -101.7864
26.8252 | -101.7875
26.8288 | -101.7891
26.8325 | -101.7911
26.8362 | -101.7936
26.8398 | -101.7967
26.8435 | -101.7992
26.8471 | -101.7998
26.8508 | -101.8011
26.8545 | -101.8014
26.8581 | -101.798
26.8618 | -101.7973
26.8654 | -101.7969
26.8691 | -101.7953
26.8727 | -101.7951
26.8764 | -101.7946
26.8801 | -101.7934
26.8837 | -101.7925
26.8874 | -101.7915
26.891 | -101.7899
26.8947 | -101.7897
26.8983 | -101.7896
26.902 | -101.7888
26.9057 | -101.7874
26.9093 | -101.7856
26.913 | -101.7826
26.9166 | -101.7797
26.9203 | -101.7773
26.924 | -101.776
26.9276 | -101.7744
26.9313 | -101.7728
26.9349 | -101.7729
26.9386 | -101.7707
26.9422 | -101.7673
26.9459 | -101.7661
26.9496 | -101.7653
26.9532 | -101.7647
26.9569 | -101.7649
26.9605 | -101.7636
26.9642 | -101.7629
26.9679 | -101.762
26.9715 | -101.7593
26.9752 | -101.7578
26.9788 | -101.7571
26.9825 | -101.7564
26.9861 | -101.756
26.9898 | -101.7558
26.9935 | -101.7562
26.9971 | -101.7566
27.0008 | -101.7566
27.0044 | -101.7572
27.0081 | -101.7578
27.0118 | -101.7577
27.0154 | -101.7586
27.0191 | -101.7593
27.0227 | -101.7593
27.0264 | -101.7588
27.03 | -101.7596
27.0337 | -101.7606
27.0374 | -101.7612
27.041 | -101.7622
27.0447 | -101.763
27.0483 | -101.7642
27.052 | -101.7667
27.0556 | -101.7692
27.0593 | -101.7706
27.063 | -101.7702
27.0666 | -101.771
27.0703 | -101.7729
27.0739 | -101.7759
27.0776 | -101.7779
27.0813 | -101.7788
27.0849 | -101.7816
27.0886 | -101.7851
27.0922 | -101.7877
27.0959 | -101.7909
27.0995 | -101.7931
27.1032 | -101.7949
27.1069 | -101.7962
27.1105 | -101.7984
27.1142 | -101.8009
27.1178 | -101.8049
27.1215 | -101.8091
27.1252 | -101.8103
27.1288 | -101.8124
27.1325 | -101.817
27.1361 | -101.8192
27.1398 | -101.8206
27.1434 | -101.8227
27.1471 | -101.8245
27.1508 | -101.828
27.1544 | -101.8299
27.1581 | -101.8323
27.1617 | -101.8343
27.1654 | -101.8357
27.1691 | -101.8373
27.1727 | -101.8394
27.1764 | -101.8413
27.18 | -101.8427
27.1837 | -101.8431
27.1873 | -101.8436
27.191 | -101.845
27.1947 | -101.8469
27.1983 | -101.8491
27.202 | -101.8513
27.2056 | -101.8517
27.2093 | -101.8521
27.2129 | -101.8522
27.2166 | -101.8513
27.2203 | -101.85
27.2239 | -101.8488
27.2276 | -101.8499
27.2312 | -101.8501
27.2349 | -101.8495
27.2386 | -101.8496
27.2422 | -101.8479
27.2459 | -101.847
27.2495 | -101.8474
27.2532 | -101.8453
27.2568 | -101.8435
27.2605 | -101.8425
27.2642 | -101.8408
27.2678 | -101.8392
27.2715 | -101.8389
27.2751 | -101.8387
27.2788 | -101.8363
27.2825 | -101.8344
27.2861 | -101.8333
27.2898 | -101.8321
27.2934 | -101.8305
27.2971 | -101.8288
27.3007 | -101.8257
27.3044 | -101.8238
27.3081 | -101.8228
27.3117 | -101.8213
27.3154 | -101.8193
27.319 | -101.8156
27.3227 | -101.8121
27.3264 | -101.8093
27.33 | -101.8056
27.3337 | -101.8004
27.3373 | -101.7964
27.341 | -101.7929
27.3446 | -101.7903
27.3483 | -101.7875
27.352 | -101.7833
27.3556 | -101.7804
27.3593 | -101.7772
27.3629 | -101.7732
27.3666 | -101.7698
27.3702 | -101.7668
27.3739 | -101.7649
27.3776 | -101.7623
27.3812 | -101.7603
27.3849 | -101.757
27.3885 | -101.7534
27.3922 | -101.7518
27.3959 | -101.7502
27.3995 | -101.7468
27.4032 | -101.7433
27.4068 | -101.7414
27.4105 | -101.7403
27.4141 | -101.7381
27.4178 | -101.7342
27.4215 | -101.7319
27.4251 | -101.7303
27.4288 | -101.7277
27.4324 | -101.7252
27.4361 | -101.7218
27.4398 | -101.7228
27.4434 | -101.7217
27.4471 | -101.7203
27.4507 | -101.7204
27.4544 | -101.7199
27.458 | -101.7187
27.4617 | -101.7179
27.4654 | -101.7174
27.469 | -101.7173
27.4727 | -101.7146
27.4763 | -101.7139
27.48 | -101.7147
27.4837 | -101.7159
27.4873 | -101.7181
27.491 | -101.7177
27.4946 | -101.7194
27.4983 | -101.7204
27.5019 | -101.7206
27.5056 | -101.719
27.5093 | -101.7157
27.5129 | -101.7158
27.5166 | -101.7183
27.5202 | -101.7206
27.5239 | -101.7226
27.5276 | -101.7235
27.5312 | -101.724
27.5349 | -101.7241
27.5385 | -101.7236
27.5422 | -101.7249
27.5458 | -101.7255
27.5495 | -101.7255
27.5532 | -101.7259
27.5568 | -101.726
27.5605 | -101.7254
27.5641 | -101.724
27.5678 | -101.7237
27.5714 | -101.7217
27.5751 | -101.7167
27.5788 | -101.7131
27.5824 | -101.7109
27.5861 | -101.7078
27.5897 | -101.705
27.5934 | -101.703
27.5971 | -101.6997
27.6007 | -101.6967
27.6044 | -101.6931
27.608 | -101.6895
27.6117 | -101.6856
27.6153 | -101.6804
27.619 | -101.6748
27.6227 | -101.6683
27.6263 | -101.6613
27.63 | -101.6551
27.6336 | -101.6489
27.6373 | -101.6419
27.641 | -101.6354
27.6446 | -101.6289
27.6483 | -101.6196
27.6519 | -101.6116
27.6556 | -101.6043
27.6592 | -101.5958
27.6629 | -101.5876
27.6666 | -101.5797
27.6702 | -101.572
27.6739 | -101.5633
27.6775 | -101.5538
27.6812 | -101.545
27.6849 | -101.5365
27.6885 | -101.5296
27.6922 | -101.522
27.6958 | -101.5141
27.6995 | -101.5071
27.7031 | -101.4981
27.7068 | -101.4908
27.7105 | -101.4857
27.7141 | -101.4799
27.7178 | -101.4737
27.7214 | -101.4682
27.7251 | -101.4622
27.7287 | -101.4554
27.7324 | -101.4485
27.7361 | -101.4436
27.7397 | -101.4401
27.7434 | -101.4363
27.747 | -101.4331
27.7507 | -101.4295
27.7544 | -101.4268
27.758 | -101.425
27.7617 | -101.4219
27.7653 | -101.4201
27.769 | -101.419
27.7726 | -101.4193
27.7763 | -101.4205
27.78 | -101.4211
27.7836 | -101.4228
27.7873 | -101.4254
27.7909 | -101.4278
27.7946 | -101.429
27.7983 | -101.4311
27.8019 | -101.4335
27.8056 | -101.4382
27.8092 | -101.4429
27.8129 | -101.4477
27.8165 | -101.4528
27.8202 | -101.4591
27.8239 | -101.4659
27.8275 | -101.4717
27.8312 | -101.4768
27.8348 | -101.481
27.8385 | -101.4852
27.8422 | -101.4917
27.8458 | -101.4991
27.8495 | -101.5065
27.8531 | -101.5112
27.8568 | -101.5124
27.8604 | -101.5099
27.8641 | -101.5121
27.8678 | -101.5159
27.8714 | -101.5195
27.8751 | -101.5233
27.8787 | -101.5263
27.8824 | -101.5298
27.886 | -101.5312
27.8897 | -101.5322
27.8934 | -101.5336
27.897 | -101.535
27.9007 | -101.5368
27.9043 | -101.5373
27.908 | -101.5351
27.9117 | -101.5317
27.9153 | -101.5265
27.919 | -101.5225
27.9226 | -101.5159
27.9263 | -101.4991
27.9299 | -101.457
27.9336 | -101.4265
27.9373 | -101.4062
27.9409 | -101.3886
27.9446 | -101.3704
27.9482 | -101.3507
27.9519 | -101.328
27.9556 | -101.3054
27.9592 | -101.2817
27.9629 | -101.2563
27.9665 | -101.2308
27.9702 | -101.1911
27.9738 | -101.1632
27.9775 | -101.133
27.9812 | -101.1017
27.9848 | -101.0687
27.9885 | -101.0353
27.9921 | -101.0005
27.9958 | -100.9555
27.9995 | -100.912
28.0031 | -100.8723