## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 273 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 276
Hood | 02-92B/P | 216
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1009.8 | hPa |  | PRC PRO 01
Humidity | 51.5 | %/Hr
Internal temperature of calibration bench | 24.29 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1017 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.8 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.98 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.96 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3104 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 17/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 21/08/2018
Technical Manager | Anthony HUE | Approval |  | 30/08/2018
CEO | Robert GLEMOT | Approval |  | 30/08/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5056 | 57.5928
0.012209 | -36.0774 | 38.4512
0.018313 | -35.0107 | 28.2591
0.024417 | -34.5474 | 22.0104
0.030521 | -34.4338 | 18.2288
0.036626 | -34.2887 | 15.1565
0.04273 | -34.1935 | 12.7901
0.048834 | -34.1476 | 11.2688
0.054938 | -34.0842 | 10.1242
0.061043 | -34.0247 | 9.0322
0.067147 | -34.0134 | 8.2044
0.073251 | -34.0134 | 7.4418
0.079355 | -34.0308 | 6.7824
0.08546 | -34.0445 | 6.1924
0.091564 | -34.0405 | 5.7525
0.097668 | -34.0235 | 5.4991
0.10377 | -34.0024 | 5.2468
0.10988 | -33.9858 | 4.9659
0.11598 | -33.9768 | 4.6709
0.12209 | -33.9782 | 4.3866
0.12819 | -33.9876 | 4.1055
0.13429 | -33.9983 | 3.8229
0.1404 | -34.0016 | 3.5491
0.1465 | -34.0036 | 3.2941
0.15261 | -34.001 | 3.0744
0.15871 | -33.9949 | 2.9106
0.16482 | -33.9872 | 2.8051
0.17092 | -33.9851 | 2.7632
0.17702 | -33.983 | 2.6891
0.18313 | -33.9822 | 2.6097
0.18923 | -33.9818 | 2.5009
0.19534 | -33.9807 | 2.3748
0.20144 | -33.9795 | 2.234
0.20754 | -33.9768 | 2.1139
0.21365 | -33.976 | 1.9999
0.21975 | -33.9756 | 1.9021
0.22586 | -33.9766 | 1.8147
0.23196 | -33.9791 | 1.7419
0.23807 | -33.9831 | 1.6785
0.24417 | -33.9872 | 1.6195
0.25027 | -33.9911 | 1.5568
0.25638 | -33.9935 | 1.4949
0.26248 | -33.9954 | 1.4284
0.26859 | -33.9965 | 1.3635
0.27469 | -33.9969 | 1.2931
0.2808 | -33.998 | 1.2205
0.2869 | -33.9989 | 1.1504
0.293 | -33.9999 | 1.0842
0.29911 | -34.0008 | 1.0246
0.30521 | -34.0018 | 0.96612
0.31132 | -34.0026 | 0.92703
0.31742 | -34.0038 | 0.88398
0.32353 | -34.0042 | 0.8484
0.32963 | -34.0051 | 0.81846
0.33573 | -34.0061 | 0.79682
0.34184 | -34.0075 | 0.77421
0.34794 | -34.0097 | 0.74943
0.35405 | -34.0118 | 0.72443
0.36015 | -34.0143 | 0.69827
0.36626 | -34.0163 | 0.67016
0.37236 | -34.0184 | 0.63324
0.37846 | -34.0206 | 0.58924
0.38457 | -34.0215 | 0.53317
0.39067 | -34.0225 | 0.48707
0.39678 | -34.0232 | 0.43593
0.40288 | -34.0237 | 0.38276
0.40899 | -34.0236 | 0.33176
0.41509 | -34.023 | 0.28176
0.42119 | -34.023 | 0.23906
0.4273 | -34.0231 | 0.19354
0.4334 | -34.023 | 0.1481
0.43951 | -34.0234 | 0.10531
0.44561 | -34.0242 | 0.065996
0.45172 | -34.0258 | 0.034483
0.45782 | -34.0266 | -0.0022716
0.46392 | -34.0278 | -0.038117
0.47003 | -34.0289 | -0.070326
0.47613 | -34.0298 | -0.10416
0.48224 | -34.0303 | -0.13866
0.48834 | -34.0294 | -0.17319
0.49445 | -34.0293 | -0.20818
0.50055 | -34.0288 | -0.24017
0.50665 | -34.0281 | -0.27419
0.51276 | -34.0271 | -0.30805
0.51886 | -34.0262 | -0.34185
0.52497 | -34.025 | -0.36628
0.53107 | -34.0243 | -0.39577
0.53717 | -34.0237 | -0.4239
0.54328 | -34.0229 | -0.44698
0.54938 | -34.0228 | -0.46967
0.55549 | -34.0228 | -0.49087
0.56159 | -34.0239 | -0.503
0.5677 | -34.0242 | -0.51807
0.5738 | -34.0253 | -0.53477
0.5799 | -34.0261 | -0.55011
0.58601 | -34.0272 | -0.56079
0.59211 | -34.0289 | -0.56785
0.59822 | -34.0302 | -0.57796
0.60432 | -34.0317 | -0.59053
0.61043 | -34.033 | -0.59913
0.61653 | -34.0342 | -0.60996
0.62263 | -34.0358 | -0.61985
0.62874 | -34.0365 | -0.63087
0.63484 | -34.0373 | -0.64062
0.64095 | -34.0379 | -0.65121
0.64705 | -34.0383 | -0.66137
0.65316 | -34.0385 | -0.67411
0.65926 | -34.0387 | -0.69005
0.66536 | -34.0376 | -0.70747
0.67147 | -34.0373 | -0.72082
0.67757 | -34.037 | -0.73523
0.68368 | -34.0363 | -0.75407
0.68978 | -34.0356 | -0.7712
0.69589 | -34.0348 | -0.79014
0.70199 | -34.0339 | -0.8126
0.70809 | -34.0331 | -0.83166
0.7142 | -34.0324 | -0.852
0.7203 | -34.0314 | -0.87109
0.72641 | -34.0309 | -0.89338
0.73251 | -34.0302 | -0.9133
0.73862 | -34.0305 | -0.93497
0.74472 | -34.0302 | -0.95494
0.75082 | -34.03 | -0.97693
0.75693 | -34.0304 | -0.99866
0.76303 | -34.0309 | -1.0205
0.76914 | -34.0327 | -1.041
0.77524 | -34.0337 | -1.0607
0.78135 | -34.035 | -1.0808
0.78745 | -34.0363 | -1.1007
0.79355 | -34.0375 | -1.1205
0.79966 | -34.0393 | -1.14
0.80576 | -34.0415 | -1.1597
0.81187 | -34.0431 | -1.1812
0.81797 | -34.045 | -1.2024
0.82408 | -34.0466 | -1.2229
0.83018 | -34.0483 | -1.2435
0.83628 | -34.0502 | -1.262
0.84239 | -34.0513 | -1.2823
0.84849 | -34.0528 | -1.3039
0.8546 | -34.0542 | -1.3236
0.8607 | -34.0554 | -1.3442
0.86681 | -34.0565 | -1.3631
0.87291 | -34.0575 | -1.3818
0.87901 | -34.0579 | -1.3996
0.88512 | -34.0587 | -1.4177
0.89122 | -34.0594 | -1.4358
0.89733 | -34.0598 | -1.4514
0.90343 | -34.0602 | -1.4689
0.90953 | -34.0604 | -1.4832
0.91564 | -34.0604 | -1.499
0.92174 | -34.0607 | -1.5139
0.92785 | -34.0608 | -1.5284
0.93395 | -34.0609 | -1.5422
0.94006 | -34.0611 | -1.5551
0.94616 | -34.0608 | -1.567
0.95226 | -34.061 | -1.5796
0.95837 | -34.0611 | -1.5934
0.96447 | -34.0613 | -1.6062
0.97058 | -34.0611 | -1.6195
0.97668 | -34.0609 | -1.6328
0.98279 | -34.0607 | -1.6456
0.98889 | -34.0605 | -1.6563
0.99499 | -34.0605 | -1.6697
1.0011 | -34.0604 | -1.6804
1.0072 | -34.0601 | -1.6921
1.0133 | -34.0598 | -1.7029
1.0194 | -34.0594 | -1.7159
1.0255 | -34.0591 | -1.7296
1.0316 | -34.0586 | -1.7415
1.0377 | -34.0582 | -1.7552
1.0438 | -34.0577 | -1.7679
1.0499 | -34.057 | -1.7832
1.056 | -34.0564 | -1.7974
1.0621 | -34.0558 | -1.8117
1.0682 | -34.0554 | -1.8242
1.0743 | -34.0548 | -1.8396
1.0805 | -34.0543 | -1.854
1.0866 | -34.0535 | -1.8691
1.0927 | -34.0527 | -1.8837
1.0988 | -34.0521 | -1.8985
1.1049 | -34.0514 | -1.9127
1.111 | -34.0506 | -1.9281
1.1171 | -34.05 | -1.9427
1.1232 | -34.0495 | -1.9579
1.1293 | -34.0492 | -1.972
1.1354 | -34.0489 | -1.987
1.1415 | -34.0486 | -2.0025
1.1476 | -34.048 | -2.0179
1.1537 | -34.0475 | -2.0326
1.1598 | -34.0476 | -2.0475
1.1659 | -34.0474 | -2.0609
1.172 | -34.0472 | -2.0748
1.1781 | -34.047 | -2.0888
1.1842 | -34.0469 | -2.103
1.1903 | -34.0472 | -2.1141
1.1964 | -34.0471 | -2.1275
1.2025 | -34.0472 | -2.1415
1.2086 | -34.0475 | -2.154
1.2147 | -34.0476 | -2.1666
1.2209 | -34.0479 | -2.1779
1.227 | -34.0486 | -2.1887
1.2331 | -34.049 | -2.2008
1.2392 | -34.0494 | -2.2126
1.2453 | -34.0498 | -2.2237
1.2514 | -34.0502 | -2.2338
1.2575 | -34.0507 | -2.2447
1.2636 | -34.0515 | -2.2536
1.2697 | -34.0521 | -2.263
1.2758 | -34.0527 | -2.2741
1.2819 | -34.0534 | -2.284
1.288 | -34.0541 | -2.2948
1.2941 | -34.0549 | -2.3033
1.3002 | -34.0555 | -2.3114
1.3063 | -34.0561 | -2.3205
1.3124 | -34.0567 | -2.3295
1.3185 | -34.0574 | -2.3385
1.3246 | -34.0581 | -2.3476
1.3307 | -34.0589 | -2.3562
1.3368 | -34.0597 | -2.3649
1.3429 | -34.0603 | -2.3738
1.349 | -34.0608 | -2.3831
1.3551 | -34.0615 | -2.3924
1.3613 | -34.062 | -2.4006
1.3674 | -34.0627 | -2.4106
1.3735 | -34.0631 | -2.4203
1.3796 | -34.0636 | -2.4293
1.3857 | -34.0639 | -2.4385
1.3918 | -34.0644 | -2.4487
1.3979 | -34.065 | -2.4579
1.404 | -34.0655 | -2.4691
1.4101 | -34.0659 | -2.4795
1.4162 | -34.0663 | -2.49
1.4223 | -34.0668 | -2.5014
1.4284 | -34.0673 | -2.5114
1.4345 | -34.0676 | -2.5228
1.4406 | -34.0679 | -2.5363
1.4467 | -34.0684 | -2.5478
1.4528 | -34.0688 | -2.5592
1.4589 | -34.0693 | -2.571
1.465 | -34.0696 | -2.5819
1.4711 | -34.0701 | -2.5933
1.4772 | -34.0706 | -2.6066
1.4833 | -34.0711 | -2.6195
1.4894 | -34.0716 | -2.6317
1.4955 | -34.0722 | -2.6437
1.5016 | -34.0727 | -2.6562
1.5078 | -34.0731 | -2.67
1.5139 | -34.0735 | -2.6831
1.52 | -34.074 | -2.6957
1.5261 | -34.0743 | -2.7083
1.5322 | -34.0746 | -2.7223
1.5383 | -34.075 | -2.7349
1.5444 | -34.0756 | -2.7476
1.5505 | -34.076 | -2.761
1.5566 | -34.0766 | -2.7741
1.5627 | -34.077 | -2.7878
1.5688 | -34.0776 | -2.8014
1.5749 | -34.078 | -2.8141
1.581 | -34.0783 | -2.8276
1.5871 | -34.0787 | -2.8413
1.5932 | -34.0791 | -2.8543
1.5993 | -34.0795 | -2.8685
1.6054 | -34.0798 | -2.8818
1.6115 | -34.0803 | -2.8959
1.6176 | -34.0806 | -2.9086
1.6237 | -34.0808 | -2.9223
1.6298 | -34.0811 | -2.935
1.6359 | -34.0814 | -2.9478
1.642 | -34.0818 | -2.96
1.6482 | -34.082 | -2.9731
1.6543 | -34.0824 | -2.9869
1.6604 | -34.0827 | -2.999
1.6665 | -34.0828 | -3.0116
1.6726 | -34.083 | -3.0243
1.6787 | -34.0834 | -3.037
1.6848 | -34.0833 | -3.0498
1.6909 | -34.0836 | -3.0627
1.697 | -34.0838 | -3.075
1.7031 | -34.0839 | -3.0878
1.7092 | -34.0838 | -3.1002
1.7153 | -34.084 | -3.1123
1.7214 | -34.0839 | -3.1247
1.7275 | -34.0838 | -3.1374
1.7336 | -34.0838 | -3.1504
1.7397 | -34.0839 | -3.1629
1.7458 | -34.0839 | -3.1753
1.7519 | -34.084 | -3.1876
1.758 | -34.084 | -3.1997
1.7641 | -34.0842 | -3.2114
1.7702 | -34.0841 | -3.2245
1.7763 | -34.0841 | -3.2362
1.7824 | -34.0841 | -3.2487
1.7885 | -34.0841 | -3.2613
1.7947 | -34.0843 | -3.2734
1.8008 | -34.0843 | -3.2848
1.8069 | -34.0843 | -3.2966
1.813 | -34.0844 | -3.3089
1.8191 | -34.0844 | -3.3216
1.8252 | -34.0845 | -3.3344
1.8313 | -34.0846 | -3.3467
1.8374 | -34.0848 | -3.3586
1.8435 | -34.0849 | -3.3706
1.8496 | -34.0849 | -3.3836
1.8557 | -34.0851 | -3.3964
1.8618 | -34.0854 | -3.4085
1.8679 | -34.0856 | -3.4206
1.874 | -34.0857 | -3.4326
1.8801 | -34.0858 | -3.4449
1.8862 | -34.0861 | -3.4569
1.8923 | -34.0863 | -3.4691
1.8984 | -34.0864 | -3.4811
1.9045 | -34.0865 | -3.4926
1.9106 | -34.0866 | -3.5047
1.9167 | -34.0867 | -3.5169
1.9228 | -34.0868 | -3.5287
1.9289 | -34.087 | -3.5405
1.9351 | -34.0871 | -3.5513
1.9412 | -34.0874 | -3.5626
1.9473 | -34.0877 | -3.5745
1.9534 | -34.0879 | -3.5859
1.9595 | -34.0881 | -3.5976
1.9656 | -34.0883 | -3.6098
1.9717 | -34.0886 | -3.6214
1.9778 | -34.0889 | -3.6327
1.9839 | -34.0892 | -3.6441
1.99 | -34.0894 | -3.6551
1.9961 | -34.0895 | -3.6672
2.0022 | -34.0897 | -3.6785
2.0083 | -34.0898 | -3.6897
2.0144 | -34.0901 | -3.7011
2.0205 | -34.0904 | -3.7127
2.0266 | -34.0907 | -3.724
2.0327 | -34.091 | -3.7355
2.0388 | -34.0914 | -3.7469
2.0449 | -34.0916 | -3.7582
2.051 | -34.0919 | -3.7695
2.0571 | -34.0922 | -3.7808
2.0632 | -34.0924 | -3.7922
2.0693 | -34.0926 | -3.8043
2.0754 | -34.0929 | -3.8156
2.0816 | -34.0932 | -3.8277
2.0877 | -34.0934 | -3.8391
2.0938 | -34.0937 | -3.8511
2.0999 | -34.0939 | -3.8628
2.106 | -34.0941 | -3.8743
2.1121 | -34.0943 | -3.8863
2.1182 | -34.0945 | -3.897
2.1243 | -34.0948 | -3.9078
2.1304 | -34.095 | -3.9189
2.1365 | -34.0953 | -3.9305
2.1426 | -34.0955 | -3.9425
2.1487 | -34.0958 | -3.9537
2.1548 | -34.0962 | -3.9655
2.1609 | -34.0966 | -3.9769
2.167 | -34.097 | -3.988
2.1731 | -34.0973 | -4
2.1792 | -34.0975 | -4.011
2.1853 | -34.0979 | -4.0229
2.1914 | -34.0982 | -4.0342
2.1975 | -34.0985 | -4.0457
2.2036 | -34.0987 | -4.0577
2.2097 | -34.099 | -4.0691
2.2158 | -34.0993 | -4.0797
2.222 | -34.0996 | -4.0902
2.2281 | -34.0998 | -4.1016
2.2342 | -34.1001 | -4.1129
2.2403 | -34.1003 | -4.1236
2.2464 | -34.1006 | -4.1344
2.2525 | -34.1008 | -4.1457
2.2586 | -34.1011 | -4.1563
2.2647 | -34.1014 | -4.1668
2.2708 | -34.1017 | -4.178
2.2769 | -34.102 | -4.1887
2.283 | -34.1022 | -4.1998
2.2891 | -34.1024 | -4.2108
2.2952 | -34.1026 | -4.2213
2.3013 | -34.1028 | -4.2322
2.3074 | -34.1031 | -4.2427
2.3135 | -34.1033 | -4.2541
2.3196 | -34.1035 | -4.2643
2.3257 | -34.1037 | -4.2749
2.3318 | -34.1038 | -4.286
2.3379 | -34.104 | -4.2968
2.344 | -34.1041 | -4.3081
2.3501 | -34.1043 | -4.3189
2.3562 | -34.1045 | -4.3293
2.3623 | -34.1046 | -4.3397
2.3685 | -34.1047 | -4.3503
2.3746 | -34.1048 | -4.3616
2.3807 | -34.105 | -4.372
2.3868 | -34.1052 | -4.3828
2.3929 | -34.1054 | -4.3935
2.399 | -34.1055 | -4.4039
2.4051 | -34.1057 | -4.4145
2.4112 | -34.1059 | -4.4252
2.4173 | -34.1061 | -4.4358
2.4234 | -34.1061 | -4.4466
2.4295 | -34.1063 | -4.4575
2.4356 | -34.1065 | -4.4683
2.4417 | -34.1066 | -4.4795
2.4478 | -34.1067 | -4.4904
2.4539 | -34.1069 | -4.5007
2.46 | -34.107 | -4.5121
2.4661 | -34.1072 | -4.5233
2.4722 | -34.1073 | -4.5335
2.4783 | -34.1074 | -4.544
2.4844 | -34.1075 | -4.555
2.4905 | -34.1076 | -4.566
2.4966 | -34.1078 | -4.5767
2.5027 | -34.1079 | -4.5873
2.5089 | -34.108 | -4.5982
2.515 | -34.1081 | -4.609
2.5211 | -34.1082 | -4.6198
2.5272 | -34.1082 | -4.6307
2.5333 | -34.1084 | -4.6417
2.5394 | -34.1086 | -4.6525
2.5455 | -34.1087 | -4.6639
2.5516 | -34.1089 | -4.6743
2.5577 | -34.1091 | -4.6857
2.5638 | -34.1092 | -4.6975
2.5699 | -34.1093 | -4.7089
2.576 | -34.1093 | -4.7205
2.5821 | -34.1095 | -4.7317
2.5882 | -34.1096 | -4.7429
2.5943 | -34.1097 | -4.7538
2.6004 | -34.1099 | -4.7647
2.6065 | -34.11 | -4.7752
2.6126 | -34.1102 | -4.7857
2.6187 | -34.1103 | -4.7963
2.6248 | -34.1105 | -4.8071
2.6309 | -34.1107 | -4.8186
2.637 | -34.1109 | -4.8302
2.6431 | -34.1111 | -4.8408
2.6492 | -34.1112 | -4.8521
2.6554 | -34.1114 | -4.863
2.6615 | -34.1115 | -4.8741
2.6676 | -34.1118 | -4.8855
2.6737 | -34.1121 | -4.8967
2.6798 | -34.1123 | -4.9079
2.6859 | -34.1125 | -4.9192
2.692 | -34.1127 | -4.9301
2.6981 | -34.1129 | -4.9412
2.7042 | -34.1131 | -4.9524
2.7103 | -34.1133 | -4.9631
2.7164 | -34.1136 | -4.9741
2.7225 | -34.1139 | -4.9855
2.7286 | -34.1141 | -4.9966
2.7347 | -34.1144 | -5.0072
2.7408 | -34.1146 | -5.0173
2.7469 | -34.1149 | -5.0286
2.753 | -34.1152 | -5.0393
2.7591 | -34.1155 | -5.0506
2.7652 | -34.1157 | -5.0621
2.7713 | -34.1159 | -5.0729
2.7774 | -34.1162 | -5.0839
2.7835 | -34.1163 | -5.0949
2.7896 | -34.1165 | -5.1055
2.7958 | -34.1168 | -5.1163
2.8019 | -34.1171 | -5.1268
2.808 | -34.1174 | -5.1377
2.8141 | -34.1177 | -5.1489
2.8202 | -34.118 | -5.1598
2.8263 | -34.1183 | -5.1708
2.8324 | -34.1187 | -5.1822
2.8385 | -34.1191 | -5.1931
2.8446 | -34.1194 | -5.2039
2.8507 | -34.1198 | -5.2147
2.8568 | -34.1202 | -5.2253
2.8629 | -34.1204 | -5.2359
2.869 | -34.1208 | -5.247
2.8751 | -34.1211 | -5.2578
2.8812 | -34.1215 | -5.2691
2.8873 | -34.1218 | -5.28
2.8934 | -34.1221 | -5.2911
2.8995 | -34.1224 | -5.3022
2.9056 | -34.1227 | -5.3134
2.9117 | -34.1231 | -5.3245
2.9178 | -34.1234 | -5.3353
2.9239 | -34.1238 | -5.3468
2.93 | -34.1242 | -5.3583
2.9361 | -34.1245 | -5.3691
2.9423 | -34.1248 | -5.3798
2.9484 | -34.1253 | -5.3909
2.9545 | -34.1256 | -5.4019
2.9606 | -34.126 | -5.4133
2.9667 | -34.1264 | -5.4245
2.9728 | -34.1267 | -5.4358
2.9789 | -34.1271 | -5.4469
2.985 | -34.1276 | -5.4581
2.9911 | -34.1279 | -5.4694
2.9972 | -34.1284 | -5.4805
3.0033 | -34.1287 | -5.492
3.0094 | -34.1291 | -5.5031
3.0155 | -34.1295 | -5.514
3.0216 | -34.1299 | -5.5253
3.0277 | -34.1302 | -5.537
3.0338 | -34.1306 | -5.5483
3.0399 | -34.131 | -5.5597
3.046 | -34.1313 | -5.5712
3.0521 | -34.1317 | -5.5822
3.0582 | -34.132 | -5.5942
3.0643 | -34.1323 | -5.6056
3.0704 | -34.1326 | -5.6174
3.0765 | -34.1329 | -5.6285
3.0827 | -34.1332 | -5.6397
3.0888 | -34.1335 | -5.6509
3.0949 | -34.1339 | -5.6622
3.101 | -34.1342 | -5.6737
3.1071 | -34.1345 | -5.6854
3.1132 | -34.1349 | -5.6971
3.1193 | -34.1353 | -5.7085
3.1254 | -34.1356 | -5.7198
3.1315 | -34.136 | -5.7316
3.1376 | -34.1363 | -5.743
3.1437 | -34.1366 | -5.7546
3.1498 | -34.1369 | -5.7665
3.1559 | -34.1372 | -5.778
3.162 | -34.1375 | -5.7899
3.1681 | -34.1378 | -5.8018
3.1742 | -34.1381 | -5.8133
3.1803 | -34.1384 | -5.8249
3.1864 | -34.1387 | -5.8361
3.1925 | -34.139 | -5.8476
3.1986 | -34.1393 | -5.8594
3.2047 | -34.1395 | -5.8714
3.2108 | -34.1398 | -5.8824
3.2169 | -34.1401 | -5.8941
3.223 | -34.1403 | -5.9057
3.2292 | -34.1406 | -5.9174
3.2353 | -34.1408 | -5.9282
3.2414 | -34.141 | -5.9396
3.2475 | -34.1412 | -5.951
3.2536 | -34.1415 | -5.9623
3.2597 | -34.1418 | -5.9732
3.2658 | -34.142 | -5.9842
3.2719 | -34.1422 | -5.9952
3.278 | -34.1424 | -6.0067
3.2841 | -34.1427 | -6.0177
3.2902 | -34.1429 | -6.0291
3.2963 | -34.1432 | -6.0405
3.3024 | -34.1433 | -6.0515
3.3085 | -34.1435 | -6.063
3.3146 | -34.1437 | -6.0742
3.3207 | -34.1439 | -6.0856
3.3268 | -34.1442 | -6.097
3.3329 | -34.1443 | -6.1077
3.339 | -34.1445 | -6.1185
3.3451 | -34.1447 | -6.1296
3.3512 | -34.1449 | -6.1403
3.3573 | -34.1451 | -6.1515
3.3634 | -34.1453 | -6.1626
3.3696 | -34.1455 | -6.1739
3.3757 | -34.1456 | -6.1846
3.3818 | -34.1459 | -6.1955
3.3879 | -34.146 | -6.2065
3.394 | -34.1462 | -6.2179
3.4001 | -34.1464 | -6.2287
3.4062 | -34.1465 | -6.2396
3.4123 | -34.1466 | -6.2502
3.4184 | -34.1468 | -6.2614
3.4245 | -34.1471 | -6.2722
3.4306 | -34.1472 | -6.2829
3.4367 | -34.1474 | -6.2934
3.4428 | -34.1475 | -6.3038
3.4489 | -34.1477 | -6.3147
3.455 | -34.1478 | -6.3256
3.4611 | -34.148 | -6.3368
3.4672 | -34.1482 | -6.3478
3.4733 | -34.1483 | -6.3589
3.4794 | -34.1485 | -6.3697
3.4855 | -34.1487 | -6.3801
3.4916 | -34.1489 | -6.3905
3.4977 | -34.149 | -6.4018
3.5038 | -34.1492 | -6.4126
3.5099 | -34.1493 | -6.4229
3.5161 | -34.1495 | -6.4334
3.5222 | -34.1496 | -6.4438
3.5283 | -34.1497 | -6.4546
3.5344 | -34.1498 | -6.4654
3.5405 | -34.15 | -6.476
3.5466 | -34.1502 | -6.4869
3.5527 | -34.1503 | -6.4978
3.5588 | -34.1504 | -6.5084
3.5649 | -34.1505 | -6.5192
3.571 | -34.1507 | -6.5298
3.5771 | -34.1508 | -6.5409
3.5832 | -34.151 | -6.5509
3.5893 | -34.1512 | -6.5615
3.5954 | -34.1514 | -6.5724
3.6015 | -34.1515 | -6.5833
3.6076 | -34.1518 | -6.5938
3.6137 | -34.1519 | -6.604
3.6198 | -34.1522 | -6.6143
3.6259 | -34.1523 | -6.6253
3.632 | -34.1525 | -6.636
3.6381 | -34.1527 | -6.6468
3.6442 | -34.1528 | -6.6574
3.6503 | -34.1531 | -6.6676
3.6565 | -34.1533 | -6.6784
3.6626 | -34.1534 | -6.689
3.6687 | -34.1537 | -6.6998
3.6748 | -34.1539 | -6.7102
3.6809 | -34.1542 | -6.7209
3.687 | -34.1544 | -6.7319
3.6931 | -34.1546 | -6.7422
3.6992 | -34.1548 | -6.7529
3.7053 | -34.155 | -6.7635
3.7114 | -34.1552 | -6.7743
3.7175 | -34.1554 | -6.785
3.7236 | -34.1557 | -6.7952
3.7297 | -34.1559 | -6.806
3.7358 | -34.1561 | -6.8169
3.7419 | -34.1563 | -6.8278
3.748 | -34.1566 | -6.8384
3.7541 | -34.1568 | -6.8491
3.7602 | -34.157 | -6.8597
3.7663 | -34.1572 | -6.8702
3.7724 | -34.1575 | -6.881
3.7785 | -34.1578 | -6.8914
3.7846 | -34.158 | -6.902
3.7907 | -34.1583 | -6.913
3.7969 | -34.1585 | -6.9237
3.803 | -34.1588 | -6.934
3.8091 | -34.159 | -6.9446
3.8152 | -34.1593 | -6.9553
3.8213 | -34.1595 | -6.9662
3.8274 | -34.1598 | -6.9765
3.8335 | -34.1601 | -6.9874
3.8396 | -34.1604 | -6.9981
3.8457 | -34.1607 | -7.0089
3.8518 | -34.161 | -7.0195
3.8579 | -34.1614 | -7.0301
3.864 | -34.1616 | -7.041
3.8701 | -34.1619 | -7.052
3.8762 | -34.1622 | -7.0625
3.8823 | -34.1624 | -7.0734
3.8884 | -34.1627 | -7.0841
3.8945 | -34.1631 | -7.0951
3.9006 | -34.1634 | -7.1061
3.9067 | -34.1636 | -7.1173
3.9128 | -34.1639 | -7.1281
3.9189 | -34.1642 | -7.1386
3.925 | -34.1644 | -7.1495
3.9311 | -34.1647 | -7.1598
3.9372 | -34.165 | -7.1708
3.9434 | -34.1653 | -7.1815
3.9495 | -34.1656 | -7.1923
3.9556 | -34.166 | -7.2029
3.9617 | -34.1663 | -7.2132
3.9678 | -34.1666 | -7.2237
3.9739 | -34.1668 | -7.2344
3.98 | -34.1671 | -7.2452
3.9861 | -34.1674 | -7.2559
3.9922 | -34.1676 | -7.2667
3.9983 | -34.1679 | -7.277
4.0044 | -34.1682 | -7.2874
4.0105 | -34.1684 | -7.2981
4.0166 | -34.1688 | -7.3089
4.0227 | -34.169 | -7.3195
4.0288 | -34.1693 | -7.3306
4.0349 | -34.1695 | -7.3411
4.041 | -34.1698 | -7.3518
4.0471 | -34.17 | -7.3626
4.0532 | -34.1703 | -7.373
4.0593 | -34.1706 | -7.3838
4.0654 | -34.1708 | -7.3943
4.0715 | -34.1711 | -7.4054
4.0776 | -34.1714 | -7.416
4.0838 | -34.1717 | -7.4272
4.0899 | -34.172 | -7.4376
4.096 | -34.1722 | -7.4483
4.1021 | -34.1725 | -7.4588
4.1082 | -34.1727 | -7.4692
4.1143 | -34.173 | -7.4795
4.1204 | -34.1733 | -7.4899
4.1265 | -34.1735 | -7.5003
4.1326 | -34.1738 | -7.5107
4.1387 | -34.1741 | -7.5212
4.1448 | -34.1744 | -7.5319
4.1509 | -34.1746 | -7.5426
4.157 | -34.1748 | -7.5536
4.1631 | -34.1751 | -7.5644
4.1692 | -34.1753 | -7.5749
4.1753 | -34.1756 | -7.5855
4.1814 | -34.1758 | -7.5957
4.1875 | -34.1761 | -7.6066
4.1936 | -34.1763 | -7.6173
4.1997 | -34.1766 | -7.6275
4.2058 | -34.1768 | -7.6384
4.2119 | -34.177 | -7.6488
4.218 | -34.1772 | -7.6591
4.2241 | -34.1775 | -7.6695
4.2303 | -34.1778 | -7.6799
4.2364 | -34.178 | -7.6904
4.2425 | -34.1782 | -7.701
4.2486 | -34.1784 | -7.7117
4.2547 | -34.1785 | -7.722
4.2608 | -34.1787 | -7.7326
4.2669 | -34.1789 | -7.7429
4.273 | -34.1791 | -7.7535
4.2791 | -34.1794 | -7.764
4.2852 | -34.1796 | -7.7747
4.2913 | -34.1797 | -7.7849
4.2974 | -34.18 | -7.7957
4.3035 | -34.1802 | -7.8065
4.3096 | -34.1804 | -7.8172
4.3157 | -34.1806 | -7.8275
4.3218 | -34.1808 | -7.8378
4.3279 | -34.1811 | -7.8483
4.334 | -34.1812 | -7.8587
4.3401 | -34.1814 | -7.8688
4.3462 | -34.1817 | -7.8795
4.3523 | -34.1819 | -7.8895
4.3584 | -34.1821 | -7.8996
4.3645 | -34.1823 | -7.91
4.3707 | -34.1825 | -7.9201
4.3768 | -34.1827 | -7.9305
4.3829 | -34.183 | -7.9412
4.389 | -34.1832 | -7.9516
4.3951 | -34.1834 | -7.9618
4.4012 | -34.1835 | -7.9725
4.4073 | -34.1837 | -7.9827
4.4134 | -34.1839 | -7.9931
4.4195 | -34.1842 | -8.0034
4.4256 | -34.1845 | -8.0136
4.4317 | -34.1847 | -8.0239
4.4378 | -34.1848 | -8.034
4.4439 | -34.185 | -8.044
4.45 | -34.1852 | -8.0544
4.4561 | -34.1853 | -8.0647
4.4622 | -34.1855 | -8.075
4.4683 | -34.1858 | -8.0854
4.4744 | -34.186 | -8.096
4.4805 | -34.1862 | -8.106
4.4866 | -34.1864 | -8.1163
4.4927 | -34.1866 | -8.1268
4.4988 | -34.1868 | -8.1371
4.5049 | -34.187 | -8.147
4.511 | -34.1873 | -8.1571
4.5172 | -34.1875 | -8.1674
4.5233 | -34.1877 | -8.178
4.5294 | -34.1879 | -8.1882
4.5355 | -34.1881 | -8.1985
4.5416 | -34.1883 | -8.2088
4.5477 | -34.1886 | -8.219
4.5538 | -34.1888 | -8.229
4.5599 | -34.189 | -8.2392
4.566 | -34.1892 | -8.2495
4.5721 | -34.1894 | -8.2598
4.5782 | -34.1897 | -8.2703
4.5843 | -34.1899 | -8.2806
4.5904 | -34.1902 | -8.291
4.5965 | -34.1904 | -8.3011
4.6026 | -34.1906 | -8.3118
4.6087 | -34.1909 | -8.322
4.6148 | -34.1911 | -8.3326
4.6209 | -34.1913 | -8.3428
4.627 | -34.1916 | -8.3533
4.6331 | -34.1917 | -8.3635
4.6392 | -34.192 | -8.3739
4.6453 | -34.1922 | -8.3841
4.6514 | -34.1924 | -8.3945
4.6576 | -34.1926 | -8.4051
4.6637 | -34.1928 | -8.4151
4.6698 | -34.193 | -8.4252
4.6759 | -34.1933 | -8.4359
4.682 | -34.1935 | -8.4462
4.6881 | -34.1936 | -8.4564
4.6942 | -34.1939 | -8.4667
4.7003 | -34.1941 | -8.477
4.7064 | -34.1943 | -8.4872
4.7125 | -34.1946 | -8.4978
4.7186 | -34.1948 | -8.5079
4.7247 | -34.1951 | -8.5183
4.7308 | -34.1953 | -8.5286
4.7369 | -34.1955 | -8.539
4.743 | -34.1958 | -8.5497
4.7491 | -34.196 | -8.5604
4.7552 | -34.1962 | -8.5708
4.7613 | -34.1964 | -8.5813
4.7674 | -34.1966 | -8.5915
4.7735 | -34.1968 | -8.6022
4.7796 | -34.1971 | -8.6128
4.7857 | -34.1974 | -8.6233
4.7918 | -34.1976 | -8.6335
4.7979 | -34.1979 | -8.6437
4.8041 | -34.1981 | -8.6541
4.8102 | -34.1984 | -8.6642
4.8163 | -34.1986 | -8.6745
4.8224 | -34.1989 | -8.6849
4.8285 | -34.1991 | -8.695
4.8346 | -34.1994 | -8.7051
4.8407 | -34.1996 | -8.7156
4.8468 | -34.1999 | -8.7263
4.8529 | -34.2001 | -8.737
4.859 | -34.2004 | -8.7476
4.8651 | -34.2006 | -8.7578
4.8712 | -34.2009 | -8.7681
4.8773 | -34.2011 | -8.7785
4.8834 | -34.2013 | -8.7889
4.8895 | -34.2016 | -8.7991
4.8956 | -34.2018 | -8.8095
4.9017 | -34.202 | -8.8198
4.9078 | -34.2022 | -8.8305
4.9139 | -34.2025 | -8.8405
4.92 | -34.2028 | -8.8506
4.9261 | -34.2031 | -8.8607
4.9322 | -34.2034 | -8.871
4.9383 | -34.2036 | -8.8814
4.9445 | -34.2039 | -8.8918
4.9506 | -34.2041 | -8.902
4.9567 | -34.2044 | -8.9122
4.9628 | -34.2046 | -8.9226
4.9689 | -34.2049 | -8.9333
4.975 | -34.2051 | -8.944
4.9811 | -34.2054 | -8.9541
4.9872 | -34.2057 | -8.9643
4.9933 | -34.206 | -8.975
4.9994 | -34.2063 | -8.9858
5.0055 | -34.2066 | -8.9962
5.0116 | -34.2069 | -9.007
5.0177 | -34.2072 | -9.0176
5.0238 | -34.2075 | -9.0284
5.0299 | -34.2078 | -9.0391
5.036 | -34.2081 | -9.0495
5.0421 | -34.2083 | -9.0598
5.0482 | -34.2086 | -9.0702
5.0543 | -34.2089 | -9.0804
5.0604 | -34.2092 | -9.0909
5.0665 | -34.2095 | -9.101
5.0726 | -34.2098 | -9.1118
5.0787 | -34.2101 | -9.1224
5.0848 | -34.2104 | -9.1329
5.091 | -34.2107 | -9.1432
5.0971 | -34.211 | -9.1533
5.1032 | -34.2113 | -9.1641
5.1093 | -34.2116 | -9.1745
5.1154 | -34.2118 | -9.1847
5.1215 | -34.2121 | -9.195
5.1276 | -34.2124 | -9.2052
5.1337 | -34.2127 | -9.2154
5.1398 | -34.213 | -9.2256
5.1459 | -34.2133 | -9.2363
5.152 | -34.2136 | -9.2469
5.1581 | -34.2139 | -9.2575
5.1642 | -34.2142 | -9.268
5.1703 | -34.2145 | -9.2783
5.1764 | -34.2148 | -9.2886
5.1825 | -34.2151 | -9.2988
5.1886 | -34.2154 | -9.3089
5.1947 | -34.2156 | -9.3193
5.2008 | -34.216 | -9.3299
5.2069 | -34.2163 | -9.3404
5.213 | -34.2166 | -9.3506
5.2191 | -34.2169 | -9.3611
5.2252 | -34.2172 | -9.3716
5.2314 | -34.2174 | -9.3823
5.2375 | -34.2177 | -9.3924
5.2436 | -34.218 | -9.4029
5.2497 | -34.2183 | -9.413
5.2558 | -34.2186 | -9.4238
5.2619 | -34.2189 | -9.4345
5.268 | -34.2193 | -9.4451
5.2741 | -34.2196 | -9.4555
5.2802 | -34.2199 | -9.4663
5.2863 | -34.2202 | -9.4767
5.2924 | -34.2206 | -9.4873
5.2985 | -34.2209 | -9.4978
5.3046 | -34.2212 | -9.5083
5.3107 | -34.2215 | -9.5189
5.3168 | -34.2218 | -9.5294
5.3229 | -34.2221 | -9.5399
5.329 | -34.2224 | -9.5503
5.3351 | -34.2227 | -9.5605
5.3412 | -34.223 | -9.5709
5.3473 | -34.2233 | -9.5814
5.3534 | -34.2236 | -9.592
5.3595 | -34.2239 | -9.6024
5.3656 | -34.2243 | -9.6129
5.3717 | -34.2246 | -9.6236
5.3779 | -34.2249 | -9.6344
5.384 | -34.2252 | -9.6448
5.3901 | -34.2255 | -9.6556
5.3962 | -34.2258 | -9.6658
5.4023 | -34.2261 | -9.6765
5.4084 | -34.2265 | -9.6869
5.4145 | -34.2268 | -9.6971
5.4206 | -34.2271 | -9.7076
5.4267 | -34.2274 | -9.7184
5.4328 | -34.2277 | -9.7289
5.4389 | -34.228 | -9.7391
5.445 | -34.2283 | -9.7499
5.4511 | -34.2287 | -9.7602
5.4572 | -34.229 | -9.7707
5.4633 | -34.2293 | -9.7812
5.4694 | -34.2296 | -9.7917
5.4755 | -34.2299 | -9.802
5.4816 | -34.2302 | -9.8125
5.4877 | -34.2305 | -9.8228
5.4938 | -34.2308 | -9.833
5.4999 | -34.2312 | -9.8434
5.506 | -34.2315 | -9.8538
5.5121 | -34.2318 | -9.8641
5.5183 | -34.2321 | -9.8743
5.5244 | -34.2324 | -9.885
5.5305 | -34.2327 | -9.8953
5.5366 | -34.233 | -9.9055
5.5427 | -34.2333 | -9.9161
5.5488 | -34.2336 | -9.9266
5.5549 | -34.2339 | -9.937
5.561 | -34.2342 | -9.9477
5.5671 | -34.2346 | -9.9579
5.5732 | -34.2349 | -9.9686
5.5793 | -34.2352 | -9.9793
5.5854 | -34.2356 | -9.9897
5.5915 | -34.2358 | -10.0004
5.5976 | -34.2361 | -10.0109
5.6037 | -34.2364 | -10.0214
5.6098 | -34.2367 | -10.0318
5.6159 | -34.2371 | -10.0421
5.622 | -34.2374 | -10.0526
5.6281 | -34.2377 | -10.0631
5.6342 | -34.238 | -10.0738
5.6403 | -34.2383 | -10.0843
5.6464 | -34.2386 | -10.0949
5.6525 | -34.2389 | -10.1055
5.6586 | -34.2392 | -10.1156
5.6648 | -34.2396 | -10.1261
5.6709 | -34.2399 | -10.1368
5.677 | -34.2402 | -10.1472
5.6831 | -34.2405 | -10.1578
5.6892 | -34.2408 | -10.1684
5.6953 | -34.2411 | -10.1788
5.7014 | -34.2414 | -10.1893
5.7075 | -34.2417 | -10.1997
5.7136 | -34.242 | -10.2101
5.7197 | -34.2423 | -10.2206
5.7258 | -34.2426 | -10.2308
5.7319 | -34.2429 | -10.2413
5.738 | -34.2432 | -10.2517
5.7441 | -34.2436 | -10.2621
5.7502 | -34.2439 | -10.2723
5.7563 | -34.2442 | -10.2828
5.7624 | -34.2445 | -10.2934
5.7685 | -34.2448 | -10.3039
5.7746 | -34.2451 | -10.3145
5.7807 | -34.2454 | -10.325
5.7868 | -34.2457 | -10.3357
5.7929 | -34.246 | -10.346
5.799 | -34.2463 | -10.3563
5.8052 | -34.2466 | -10.3666
5.8113 | -34.2469 | -10.3769
5.8174 | -34.2472 | -10.3872
5.8235 | -34.2475 | -10.3976
5.8296 | -34.2478 | -10.4082
5.8357 | -34.248 | -10.4185
5.8418 | -34.2484 | -10.4291
5.8479 | -34.2486 | -10.4396
5.854 | -34.2489 | -10.4501
5.8601 | -34.2493 | -10.4605
5.8662 | -34.2496 | -10.4711
5.8723 | -34.2498 | -10.4813
5.8784 | -34.2501 | -10.4916
5.8845 | -34.2504 | -10.5019
5.8906 | -34.2507 | -10.5124
5.8967 | -34.251 | -10.5227
5.9028 | -34.2514 | -10.5329
5.9089 | -34.2517 | -10.5432
5.915 | -34.252 | -10.5534
5.9211 | -34.2523 | -10.5637
5.9272 | -34.2526 | -10.5739
5.9333 | -34.2529 | -10.5841
5.9394 | -34.2532 | -10.5944
5.9455 | -34.2536 | -10.6046
5.9517 | -34.2539 | -10.6148
5.9578 | -34.2543 | -10.6253
5.9639 | -34.2546 | -10.6354
5.97 | -34.2549 | -10.6459
5.9761 | -34.2553 | -10.6561
5.9822 | -34.2555 | -10.6666
5.9883 | -34.2559 | -10.677
5.9944 | -34.2562 | -10.6874
6.0005 | -34.2565 | -10.6979
6.0066 | -34.2568 | -10.7081
6.0127 | -34.2572 | -10.7182
6.0188 | -34.2575 | -10.7288
6.0249 | -34.2578 | -10.7392
6.031 | -34.2581 | -10.7493
6.0371 | -34.2584 | -10.7597
6.0432 | -34.2587 | -10.7699
6.0493 | -34.259 | -10.7802
6.0554 | -34.2594 | -10.7903
6.0615 | -34.2597 | -10.8007
6.0676 | -34.26 | -10.8112
6.0737 | -34.2603 | -10.8216
6.0798 | -34.2606 | -10.832
6.0859 | -34.261 | -10.8422
6.0921 | -34.2613 | -10.8524
6.0982 | -34.2616 | -10.8627
6.1043 | -34.2619 | -10.8729
6.1104 | -34.2623 | -10.8834
6.1165 | -34.2626 | -10.8933
6.1226 | -34.2629 | -10.9035
6.1287 | -34.2632 | -10.9138
6.1348 | -34.2636 | -10.924
6.1409 | -34.2639 | -10.9343
6.147 | -34.2642 | -10.9445
6.1531 | -34.2645 | -10.9548
6.1592 | -34.2648 | -10.965
6.1653 | -34.2651 | -10.9751
6.1714 | -34.2654 | -10.9855
6.1775 | -34.2658 | -10.9959
6.1836 | -34.2661 | -11.0062
6.1897 | -34.2664 | -11.0164
6.1958 | -34.2667 | -11.027
6.2019 | -34.2671 | -11.0373
6.208 | -34.2674 | -11.0476
6.2141 | -34.2677 | -11.0578
6.2202 | -34.2681 | -11.0682
6.2263 | -34.2684 | -11.0785
6.2325 | -34.2687 | -11.0885
6.2386 | -34.2691 | -11.0985
6.2447 | -34.2694 | -11.109
6.2508 | -34.2697 | -11.1192
6.2569 | -34.2701 | -11.1294
6.263 | -34.2704 | -11.1397
6.2691 | -34.2707 | -11.1499
6.2752 | -34.271 | -11.1602
6.2813 | -34.2713 | -11.1704
6.2874 | -34.2716 | -11.1806
6.2935 | -34.2719 | -11.1908
6.2996 | -34.2722 | -11.2009
6.3057 | -34.2726 | -11.2113
6.3118 | -34.2728 | -11.2216
6.3179 | -34.2732 | -11.2319
6.324 | -34.2735 | -11.2421
6.3301 | -34.2738 | -11.2524
6.3362 | -34.2741 | -11.2625
6.3423 | -34.2745 | -11.2729
6.3484 | -34.2748 | -11.2833
6.3545 | -34.2751 | -11.2933
6.3606 | -34.2754 | -11.3032
6.3667 | -34.2757 | -11.3134
6.3728 | -34.276 | -11.3236
6.379 | -34.2764 | -11.3339
6.3851 | -34.2767 | -11.3439
6.3912 | -34.277 | -11.3541
6.3973 | -34.2773 | -11.3643
6.4034 | -34.2776 | -11.3744
6.4095 | -34.2779 | -11.3846
6.4156 | -34.2782 | -11.3949
6.4217 | -34.2785 | -11.405
6.4278 | -34.2789 | -11.4153
6.4339 | -34.2792 | -11.4254
6.44 | -34.2796 | -11.4355
6.4461 | -34.2798 | -11.4459
6.4522 | -34.2802 | -11.4559
6.4583 | -34.2804 | -11.4662
6.4644 | -34.2808 | -11.4765
6.4705 | -34.2811 | -11.4865
6.4766 | -34.2814 | -11.4967
6.4827 | -34.2817 | -11.5069
6.4888 | -34.282 | -11.517
6.4949 | -34.2823 | -11.5269
6.501 | -34.2826 | -11.5372
6.5071 | -34.283 | -11.5472
6.5132 | -34.2833 | -11.5574
6.5194 | -34.2836 | -11.5677
6.5255 | -34.2839 | -11.5777
6.5316 | -34.2842 | -11.588
6.5377 | -34.2845 | -11.5982
6.5438 | -34.2848 | -11.6082
6.5499 | -34.2852 | -11.6183
6.556 | -34.2855 | -11.6284
6.5621 | -34.2858 | -11.6388
6.5682 | -34.2861 | -11.649
6.5743 | -34.2865 | -11.6591
6.5804 | -34.2868 | -11.6694
6.5865 | -34.2871 | -11.6796
6.5926 | -34.2874 | -11.6897
6.5987 | -34.2877 | -11.6996
6.6048 | -34.2881 | -11.7099
6.6109 | -34.2884 | -11.7201
6.617 | -34.2887 | -11.7304
6.6231 | -34.289 | -11.7406
6.6292 | -34.2893 | -11.7508
6.6353 | -34.2896 | -11.7612
6.6414 | -34.29 | -11.7713
6.6475 | -34.2903 | -11.7815
6.6536 | -34.2906 | -11.7917
6.6597 | -34.2909 | -11.8016
6.6659 | -34.2912 | -11.8118
6.672 | -34.2916 | -11.8218
6.6781 | -34.2919 | -11.8321
6.6842 | -34.2922 | -11.8421
6.6903 | -34.2926 | -11.8521
6.6964 | -34.2929 | -11.8623
6.7025 | -34.2932 | -11.8723
6.7086 | -34.2935 | -11.8824
6.7147 | -34.2938 | -11.8926
6.7208 | -34.2941 | -11.9029
6.7269 | -34.2944 | -11.9129
6.733 | -34.2947 | -11.9229
6.7391 | -34.2951 | -11.9331
6.7452 | -34.2954 | -11.943
6.7513 | -34.2957 | -11.9529
6.7574 | -34.2961 | -11.9631
6.7635 | -34.2964 | -11.9731
6.7696 | -34.2967 | -11.9831
6.7757 | -34.2971 | -11.9932
6.7818 | -34.2974 | -12.0034
6.7879 | -34.2977 | -12.0134
6.794 | -34.298 | -12.0233
6.8001 | -34.2983 | -12.0334
6.8063 | -34.2986 | -12.0436
6.8124 | -34.299 | -12.0538
6.8185 | -34.2993 | -12.0641
6.8246 | -34.2996 | -12.0741
6.8307 | -34.2999 | -12.0842
6.8368 | -34.3002 | -12.0945
6.8429 | -34.3005 | -12.1047
6.849 | -34.3009 | -12.1148
6.8551 | -34.3012 | -12.1248
6.8612 | -34.3015 | -12.1349
6.8673 | -34.3018 | -12.145
6.8734 | -34.3022 | -12.1553
6.8795 | -34.3025 | -12.1653
6.8856 | -34.3028 | -12.1754
6.8917 | -34.3032 | -12.1855
6.8978 | -34.3035 | -12.1954
6.9039 | -34.3038 | -12.2052
6.91 | -34.3041 | -12.2154
6.9161 | -34.3045 | -12.2253
6.9222 | -34.3048 | -12.2354
6.9283 | -34.3051 | -12.2452
6.9344 | -34.3055 | -12.2553
6.9405 | -34.3058 | -12.2653
6.9466 | -34.3062 | -12.2756
6.9528 | -34.3065 | -12.2859
6.9589 | -34.3069 | -12.2961
6.965 | -34.3072 | -12.3062
6.9711 | -34.3076 | -12.3164
6.9772 | -34.3079 | -12.3265
6.9833 | -34.3082 | -12.3365
6.9894 | -34.3086 | -12.3468
6.9955 | -34.3089 | -12.3567
7.0016 | -34.3092 | -12.3667
7.0077 | -34.3096 | -12.3768
7.0138 | -34.3099 | -12.3868
7.0199 | -34.3102 | -12.3968
7.026 | -34.3106 | -12.4068
7.0321 | -34.3109 | -12.4167
7.0382 | -34.3112 | -12.4265
7.0443 | -34.3116 | -12.4367
7.0504 | -34.3119 | -12.4468
7.0565 | -34.3123 | -12.4569
7.0626 | -34.3126 | -12.4667
7.0687 | -34.3129 | -12.4764
7.0748 | -34.3133 | -12.4862
7.0809 | -34.3136 | -12.4961
7.087 | -34.314 | -12.5063
7.0932 | -34.3143 | -12.5164
7.0993 | -34.3146 | -12.5264
7.1054 | -34.315 | -12.5363
7.1115 | -34.3153 | -12.5461
7.1176 | -34.3156 | -12.5561
7.1237 | -34.316 | -12.5661
7.1298 | -34.3163 | -12.576
7.1359 | -34.3166 | -12.586
7.142 | -34.317 | -12.5958
7.1481 | -34.3173 | -12.6056
7.1542 | -34.3176 | -12.6157
7.1603 | -34.3179 | -12.6257
7.1664 | -34.3183 | -12.6356
7.1725 | -34.3187 | -12.6456
7.1786 | -34.3191 | -12.6556
7.1847 | -34.3194 | -12.6656
7.1908 | -34.3197 | -12.6758
7.1969 | -34.32 | -12.6859
7.203 | -34.3204 | -12.6957
7.2091 | -34.3207 | -12.7056
7.2152 | -34.3211 | -12.7156
7.2213 | -34.3214 | -12.7253
7.2274 | -34.3217 | -12.7353
7.2335 | -34.3221 | -12.7452
7.2397 | -34.3224 | -12.7551
7.2458 | -34.3228 | -12.7648
7.2519 | -34.3232 | -12.7748
7.258 | -34.3235 | -12.7849
7.2641 | -34.3238 | -12.7949
7.2702 | -34.3241 | -12.805
7.2763 | -34.3245 | -12.815
7.2824 | -34.3249 | -12.825
7.2885 | -34.3252 | -12.835
7.2946 | -34.3256 | -12.8449
7.3007 | -34.3259 | -12.855
7.3068 | -34.3262 | -12.8651
7.3129 | -34.3266 | -12.8751
7.319 | -34.3269 | -12.885
7.3251 | -34.3273 | -12.8948
7.3312 | -34.3276 | -12.9048
7.3373 | -34.328 | -12.9145
7.3434 | -34.3283 | -12.9244
7.3495 | -34.3287 | -12.9341
7.3556 | -34.329 | -12.9442
7.3617 | -34.3294 | -12.9543
7.3678 | -34.3298 | -12.9642
7.3739 | -34.3301 | -12.9742
7.3801 | -34.3304 | -12.9844
7.3862 | -34.3308 | -12.9943
7.3923 | -34.3311 | -13.0044
7.3984 | -34.3315 | -13.0141
7.4045 | -34.3318 | -13.024
7.4106 | -34.3322 | -13.0338
7.4167 | -34.3325 | -13.0439
7.4228 | -34.3329 | -13.0538
7.4289 | -34.3333 | -13.0637
7.435 | -34.3336 | -13.0735
7.4411 | -34.334 | -13.0834
7.4472 | -34.3343 | -13.0934
7.4533 | -34.3347 | -13.1031
7.4594 | -34.335 | -13.1132
7.4655 | -34.3354 | -13.123
7.4716 | -34.3357 | -13.1331
7.4777 | -34.3361 | -13.143
7.4838 | -34.3364 | -13.1528
7.4899 | -34.3368 | -13.1626
7.496 | -34.3371 | -13.1723
7.5021 | -34.3375 | -13.1821
7.5082 | -34.3378 | -13.1921
7.5143 | -34.3382 | -13.202
7.5204 | -34.3385 | -13.212
7.5266 | -34.3389 | -13.2218
7.5327 | -34.3392 | -13.2316
7.5388 | -34.3395 | -13.2417
7.5449 | -34.3399 | -13.2518
7.551 | -34.3403 | -13.2616
7.5571 | -34.3406 | -13.2716
7.5632 | -34.341 | -13.2815
7.5693 | -34.3414 | -13.2913
7.5754 | -34.3417 | -13.3011
7.5815 | -34.3421 | -13.311
7.5876 | -34.3424 | -13.3211
7.5937 | -34.3428 | -13.3311
7.5998 | -34.3432 | -13.3409
7.6059 | -34.3436 | -13.351
7.612 | -34.3439 | -13.3609
7.6181 | -34.3443 | -13.3708
7.6242 | -34.3446 | -13.3809
7.6303 | -34.345 | -13.391
7.6364 | -34.3453 | -13.401
7.6425 | -34.3457 | -13.4109
7.6486 | -34.346 | -13.4209
7.6547 | -34.3464 | -13.4307
7.6608 | -34.3467 | -13.4405
7.667 | -34.3471 | -13.4505
7.6731 | -34.3474 | -13.4605
7.6792 | -34.3478 | -13.4706
7.6853 | -34.3481 | -13.4806
7.6914 | -34.3485 | -13.4906
7.6975 | -34.3489 | -13.5006
7.7036 | -34.3492 | -13.5105
7.7097 | -34.3496 | -13.5206
7.7158 | -34.35 | -13.5305
7.7219 | -34.3503 | -13.5403
7.728 | -34.3507 | -13.5503
7.7341 | -34.351 | -13.5602
7.7402 | -34.3514 | -13.5702
7.7463 | -34.3518 | -13.5802
7.7524 | -34.3521 | -13.5904
7.7585 | -34.3525 | -13.6003
7.7646 | -34.3528 | -13.6102
7.7707 | -34.3532 | -13.6203
7.7768 | -34.3535 | -13.6303
7.7829 | -34.3539 | -13.6405
7.789 | -34.3542 | -13.6507
7.7951 | -34.3546 | -13.6606
7.8012 | -34.3549 | -13.6706
7.8073 | -34.3553 | -13.6807
7.8135 | -34.3557 | -13.6906
7.8196 | -34.3561 | -13.7007
7.8257 | -34.3564 | -13.7108
7.8318 | -34.3567 | -13.7209
7.8379 | -34.3571 | -13.7307
7.844 | -34.3575 | -13.7406
7.8501 | -34.3578 | -13.7508
7.8562 | -34.3582 | -13.761
7.8623 | -34.3586 | -13.7708
7.8684 | -34.3589 | -13.7809
7.8745 | -34.3592 | -13.7909
7.8806 | -34.3596 | -13.801
7.8867 | -34.36 | -13.8107
7.8928 | -34.3603 | -13.8205
7.8989 | -34.3607 | -13.8305
7.905 | -34.3611 | -13.8406
7.9111 | -34.3614 | -13.8505
7.9172 | -34.3618 | -13.8605
7.9233 | -34.3621 | -13.8706
7.9294 | -34.3625 | -13.8806
7.9355 | -34.3628 | -13.8904
7.9416 | -34.3631 | -13.9004
7.9477 | -34.3635 | -13.9102
7.9539 | -34.3638 | -13.9201
7.96 | -34.3642 | -13.9302
7.9661 | -34.3646 | -13.9402
7.9722 | -34.3649 | -13.95
7.9783 | -34.3653 | -13.9602
7.9844 | -34.3657 | -13.97
7.9905 | -34.366 | -13.9802
7.9966 | -34.3664 | -13.9903
8.0027 | -34.3667 | -14.0002
8.0088 | -34.3671 | -14.0102
8.0149 | -34.3674 | -14.0202
8.021 | -34.3678 | -14.0302
8.0271 | -34.3681 | -14.0404
8.0332 | -34.3685 | -14.0503
8.0393 | -34.3689 | -14.0603
8.0454 | -34.3692 | -14.0705
8.0515 | -34.3696 | -14.0805
8.0576 | -34.37 | -14.0907
8.0637 | -34.3703 | -14.1006
8.0698 | -34.3707 | -14.1104
8.0759 | -34.371 | -14.1203
8.082 | -34.3714 | -14.1302
8.0881 | -34.3717 | -14.1402
8.0942 | -34.3721 | -14.1503
8.1004 | -34.3725 | -14.1603
8.1065 | -34.3728 | -14.1702
8.1126 | -34.3732 | -14.1802
8.1187 | -34.3735 | -14.1903
8.1248 | -34.3739 | -14.2001
8.1309 | -34.3743 | -14.2099
8.137 | -34.3746 | -14.2197
8.1431 | -34.375 | -14.2299
8.1492 | -34.3753 | -14.2398
8.1553 | -34.3757 | -14.2499
8.1614 | -34.3761 | -14.26
8.1675 | -34.3764 | -14.27
8.1736 | -34.3767 | -14.2801
8.1797 | -34.3771 | -14.2901
8.1858 | -34.3775 | -14.3002
8.1919 | -34.3778 | -14.3101
8.198 | -34.3782 | -14.3199
8.2041 | -34.3786 | -14.33
8.2102 | -34.3789 | -14.3399
8.2163 | -34.3793 | -14.3498
8.2224 | -34.3796 | -14.3598
8.2285 | -34.38 | -14.3695
8.2346 | -34.3803 | -14.3794
8.2408 | -34.3807 | -14.3893
8.2469 | -34.3811 | -14.3993
8.253 | -34.3814 | -14.4092
8.2591 | -34.3818 | -14.4192
8.2652 | -34.3822 | -14.4292
8.2713 | -34.3825 | -14.4392
8.2774 | -34.3829 | -14.4493
8.2835 | -34.3833 | -14.4594
8.2896 | -34.3836 | -14.4692
8.2957 | -34.384 | -14.4793
8.3018 | -34.3844 | -14.4893
8.3079 | -34.3847 | -14.4991
8.314 | -34.3851 | -14.5091
8.3201 | -34.3855 | -14.519
8.3262 | -34.3858 | -14.529
8.3323 | -34.3862 | -14.5389
8.3384 | -34.3866 | -14.549
8.3445 | -34.3869 | -14.5588
8.3506 | -34.3873 | -14.5688
8.3567 | -34.3877 | -14.5788
8.3628 | -34.388 | -14.5888
8.3689 | -34.3884 | -14.5987
8.375 | -34.3887 | -14.6086
8.3812 | -34.3891 | -14.6184
8.3873 | -34.3895 | -14.6284
8.3934 | -34.3898 | -14.6384
8.3995 | -34.3901 | -14.6482
8.4056 | -34.3905 | -14.6582
8.4117 | -34.3909 | -14.6682
8.4178 | -34.3912 | -14.6781
8.4239 | -34.3916 | -14.6879
8.43 | -34.3919 | -14.6979
8.4361 | -34.3923 | -14.708
8.4422 | -34.3927 | -14.7178
8.4483 | -34.393 | -14.7277
8.4544 | -34.3934 | -14.7376
8.4605 | -34.3938 | -14.7474
8.4666 | -34.3941 | -14.7575
8.4727 | -34.3945 | -14.7675
8.4788 | -34.3948 | -14.7774
8.4849 | -34.3952 | -14.7873
8.491 | -34.3955 | -14.7973
8.4971 | -34.3959 | -14.8071
8.5032 | -34.3963 | -14.8169
8.5093 | -34.3967 | -14.8267
8.5154 | -34.397 | -14.8366
8.5215 | -34.3974 | -14.8464
8.5277 | -34.3978 | -14.8564
8.5338 | -34.3981 | -14.8662
8.5399 | -34.3985 | -14.8761
8.546 | -34.3989 | -14.8859
8.5521 | -34.3992 | -14.8958
8.5582 | -34.3996 | -14.9057
8.5643 | -34.3999 | -14.9154
8.5704 | -34.4003 | -14.9251
8.5765 | -34.4007 | -14.9351
8.5826 | -34.4011 | -14.9447
8.5887 | -34.4015 | -14.9545
8.5948 | -34.4018 | -14.9644
8.6009 | -34.4022 | -14.9741
8.607 | -34.4026 | -14.9841
8.6131 | -34.4029 | -14.994
8.6192 | -34.4033 | -15.0037
8.6253 | -34.4037 | -15.0135
8.6314 | -34.4041 | -15.0234
8.6375 | -34.4044 | -15.0333
8.6436 | -34.4048 | -15.0431
8.6497 | -34.4052 | -15.0529
8.6558 | -34.4055 | -15.0628
8.6619 | -34.4059 | -15.0727
8.6681 | -34.4062 | -15.0827
8.6742 | -34.4066 | -15.0925
8.6803 | -34.407 | -15.1023
8.6864 | -34.4074 | -15.1119
8.6925 | -34.4078 | -15.1216
8.6986 | -34.4081 | -15.1315
8.7047 | -34.4085 | -15.1415
8.7108 | -34.4088 | -15.1513
8.7169 | -34.4092 | -15.1612
8.723 | -34.4096 | -15.1709
8.7291 | -34.41 | -15.1806
8.7352 | -34.4103 | -15.1906
8.7413 | -34.4107 | -15.2005
8.7474 | -34.411 | -15.2102
8.7535 | -34.4114 | -15.22
8.7596 | -34.4118 | -15.2299
8.7657 | -34.4122 | -15.2399
8.7718 | -34.4125 | -15.2498
8.7779 | -34.4129 | -15.2597
8.784 | -34.4133 | -15.2695
8.7901 | -34.4137 | -15.2794
8.7962 | -34.4141 | -15.289
8.8023 | -34.4144 | -15.2988
8.8084 | -34.4148 | -15.3086
8.8146 | -34.4152 | -15.3184
8.8207 | -34.4155 | -15.3282
8.8268 | -34.416 | -15.3378
8.8329 | -34.4163 | -15.3476
8.839 | -34.4167 | -15.3574
8.8451 | -34.4171 | -15.3671
8.8512 | -34.4175 | -15.3768
8.8573 | -34.4178 | -15.3867
8.8634 | -34.4182 | -15.3963
8.8695 | -34.4186 | -15.4062
8.8756 | -34.419 | -15.4161
8.8817 | -34.4193 | -15.4258
8.8878 | -34.4197 | -15.4358
8.8939 | -34.4201 | -15.4456
8.9 | -34.4205 | -15.4553
8.9061 | -34.4209 | -15.465
8.9122 | -34.4212 | -15.4749
8.9183 | -34.4216 | -15.4847
8.9244 | -34.422 | -15.4946
8.9305 | -34.4225 | -15.5042
8.9366 | -34.4228 | -15.5139
8.9427 | -34.4232 | -15.5238
8.9488 | -34.4236 | -15.5335
8.955 | -34.424 | -15.5433
8.9611 | -34.4244 | -15.5533
8.9672 | -34.4248 | -15.563
8.9733 | -34.4252 | -15.5726
8.9794 | -34.4256 | -15.5823
8.9855 | -34.426 | -15.5922
8.9916 | -34.4263 | -15.6021
8.9977 | -34.4267 | -15.6118
9.0038 | -34.4271 | -15.6215
9.0099 | -34.4275 | -15.6313
9.016 | -34.4279 | -15.6411
9.0221 | -34.4283 | -15.6507
9.0282 | -34.4287 | -15.6602
9.0343 | -34.4291 | -15.67
9.0404 | -34.4295 | -15.6797
9.0465 | -34.4299 | -15.6896
9.0526 | -34.4303 | -15.6993
9.0587 | -34.4307 | -15.709
9.0648 | -34.4311 | -15.7187
9.0709 | -34.4315 | -15.7284
9.077 | -34.4319 | -15.7383
9.0831 | -34.4323 | -15.7481
9.0892 | -34.4327 | -15.7579
9.0953 | -34.4331 | -15.7678
9.1015 | -34.4335 | -15.7775
9.1076 | -34.4339 | -15.7871
9.1137 | -34.4343 | -15.7968
9.1198 | -34.4347 | -15.8065
9.1259 | -34.4351 | -15.8162
9.132 | -34.4355 | -15.8261
9.1381 | -34.4359 | -15.8358
9.1442 | -34.4363 | -15.8452
9.1503 | -34.4367 | -15.8551
9.1564 | -34.4371 | -15.8648
9.1625 | -34.4375 | -15.8746
9.1686 | -34.4379 | -15.8843
9.1747 | -34.4383 | -15.894
9.1808 | -34.4388 | -15.9036
9.1869 | -34.4392 | -15.9134
9.193 | -34.4395 | -15.9232
9.1991 | -34.4399 | -15.9329
9.2052 | -34.4403 | -15.9427
9.2113 | -34.4408 | -15.9524
9.2174 | -34.4412 | -15.9621
9.2235 | -34.4416 | -15.9718
9.2296 | -34.442 | -15.9815
9.2357 | -34.4424 | -15.9914
9.2419 | -34.4428 | -16.0011
9.248 | -34.4433 | -16.0108
9.2541 | -34.4437 | -16.0205
9.2602 | -34.4441 | -16.0304
9.2663 | -34.4445 | -16.0399
9.2724 | -34.4449 | -16.0496
9.2785 | -34.4453 | -16.0593
9.2846 | -34.4458 | -16.0689
9.2907 | -34.4462 | -16.0786
9.2968 | -34.4466 | -16.0885
9.3029 | -34.447 | -16.0983
9.309 | -34.4474 | -16.108
9.3151 | -34.4478 | -16.1176
9.3212 | -34.4482 | -16.1272
9.3273 | -34.4487 | -16.1369
9.3334 | -34.4491 | -16.1465
9.3395 | -34.4495 | -16.1561
9.3456 | -34.4499 | -16.1656
9.3517 | -34.4503 | -16.1754
9.3578 | -34.4507 | -16.1851
9.3639 | -34.4511 | -16.1946
9.37 | -34.4515 | -16.2043
9.3761 | -34.4519 | -16.2142
9.3822 | -34.4524 | -16.2239
9.3884 | -34.4528 | -16.2334
9.3945 | -34.4532 | -16.2431
9.4006 | -34.4537 | -16.2528
9.4067 | -34.4541 | -16.2626
9.4128 | -34.4545 | -16.2723
9.4189 | -34.4549 | -16.2822
9.425 | -34.4553 | -16.2918
9.4311 | -34.4558 | -16.3014
9.4372 | -34.4562 | -16.311
9.4433 | -34.4566 | -16.3207
9.4494 | -34.457 | -16.3303
9.4555 | -34.4575 | -16.34
9.4616 | -34.4579 | -16.3496
9.4677 | -34.4583 | -16.3593
9.4738 | -34.4587 | -16.369
9.4799 | -34.4592 | -16.3787
9.486 | -34.4596 | -16.3884
9.4921 | -34.46 | -16.398
9.4982 | -34.4605 | -16.4075
9.5043 | -34.4609 | -16.4171
9.5104 | -34.4613 | -16.4267
9.5165 | -34.4617 | -16.4364
9.5226 | -34.4621 | -16.4462
9.5288 | -34.4625 | -16.4557
9.5349 | -34.463 | -16.4654
9.541 | -34.4634 | -16.4748
9.5471 | -34.4638 | -16.4845
9.5532 | -34.4642 | -16.4942
9.5593 | -34.4647 | -16.504
9.5654 | -34.4651 | -16.5136
9.5715 | -34.4655 | -16.5232
9.5776 | -34.4659 | -16.533
9.5837 | -34.4663 | -16.5425
9.5898 | -34.4667 | -16.5523
9.5959 | -34.4672 | -16.562
9.602 | -34.4676 | -16.5715
9.6081 | -34.468 | -16.5811
9.6142 | -34.4684 | -16.5907
9.6203 | -34.4688 | -16.6004
9.6264 | -34.4692 | -16.6099
9.6325 | -34.4697 | -16.6196
9.6386 | -34.4701 | -16.6291
9.6447 | -34.4705 | -16.6388
9.6508 | -34.4709 | -16.6483
9.6569 | -34.4714 | -16.658
9.663 | -34.4718 | -16.6677
9.6691 | -34.4722 | -16.6773
9.6753 | -34.4727 | -16.6868
9.6814 | -34.4731 | -16.6963
9.6875 | -34.4735 | -16.7057
9.6936 | -34.4739 | -16.7152
9.6997 | -34.4744 | -16.725
9.7058 | -34.4748 | -16.7345
9.7119 | -34.4753 | -16.7442
9.718 | -34.4757 | -16.754
9.7241 | -34.4761 | -16.7635
9.7302 | -34.4765 | -16.7731
9.7363 | -34.477 | -16.7827
9.7424 | -34.4774 | -16.7921
9.7485 | -34.4778 | -16.8018
9.7546 | -34.4782 | -16.8112
9.7607 | -34.4787 | -16.8209
9.7668 | -34.4791 | -16.8306
9.7729 | -34.4795 | -16.8401
9.779 | -34.48 | -16.8495
9.7851 | -34.4804 | -16.8593
9.7912 | -34.4809 | -16.8688
9.7973 | -34.4813 | -16.8785
9.8034 | -34.4817 | -16.8879
9.8095 | -34.4822 | -16.8977
9.8157 | -34.4826 | -16.9073
9.8218 | -34.4831 | -16.9171
9.8279 | -34.4835 | -16.9265
9.834 | -34.4839 | -16.9362
9.8401 | -34.4843 | -16.9458
9.8462 | -34.4848 | -16.9551
9.8523 | -34.4852 | -16.9647
9.8584 | -34.4856 | -16.9742
9.8645 | -34.486 | -16.984
9.8706 | -34.4864 | -16.9937
9.8767 | -34.4868 | -17.0032
9.8828 | -34.4873 | -17.0126
9.8889 | -34.4877 | -17.022
9.895 | -34.4881 | -17.0316
9.9011 | -34.4886 | -17.0412
9.9072 | -34.489 | -17.0507
9.9133 | -34.4894 | -17.0602
9.9194 | -34.4898 | -17.0698
9.9255 | -34.4902 | -17.0793
9.9316 | -34.4907 | -17.0889
9.9377 | -34.4911 | -17.0984
9.9438 | -34.4915 | -17.1081
9.9499 | -34.4919 | -17.1178
9.956 | -34.4924 | -17.1273
9.9622 | -34.4928 | -17.137
9.9683 | -34.4932 | -17.1465
9.9744 | -34.4936 | -17.1562
9.9805 | -34.4941 | -17.1657
9.9866 | -34.4945 | -17.1752
9.9927 | -34.4949 | -17.1847
9.9988 | -34.4953 | -17.1944
10.0049 | -34.4958 | -17.2039
10.011 | -34.4962 | -17.2135
10.0171 | -34.4966 | -17.223
10.0232 | -34.4971 | -17.2323
10.0293 | -34.4975 | -17.242
10.0354 | -34.498 | -17.2515
10.0415 | -34.4984 | -17.261
10.0476 | -34.4988 | -17.2707
10.0537 | -34.4993 | -17.2803
10.0598 | -34.4997 | -17.2898
10.0659 | -34.5001 | -17.2994
10.072 | -34.5005 | -17.3088
10.0781 | -34.501 | -17.3182
10.0842 | -34.5014 | -17.3277
10.0903 | -34.5019 | -17.3372
10.0964 | -34.5023 | -17.3465
10.1026 | -34.5027 | -17.3561
10.1087 | -34.5031 | -17.3656
10.1148 | -34.5035 | -17.3752
10.1209 | -34.5039 | -17.3846
10.127 | -34.5044 | -17.3939
10.1331 | -34.5048 | -17.4033
10.1392 | -34.5052 | -17.4127
10.1453 | -34.5056 | -17.4221
10.1514 | -34.5061 | -17.4316
10.1575 | -34.5065 | -17.4411
10.1636 | -34.5069 | -17.4504
10.1697 | -34.5073 | -17.46
10.1758 | -34.5078 | -17.4693
10.1819 | -34.5082 | -17.4787
10.188 | -34.5086 | -17.4882
10.1941 | -34.509 | -17.4977
10.2002 | -34.5094 | -17.507
10.2063 | -34.5099 | -17.5165
10.2124 | -34.5103 | -17.5261
10.2185 | -34.5107 | -17.5355
10.2246 | -34.5111 | -17.5449
10.2307 | -34.5116 | -17.5544
10.2368 | -34.512 | -17.5636
10.2429 | -34.5124 | -17.573
10.2491 | -34.5128 | -17.5825
10.2552 | -34.5133 | -17.5921
10.2613 | -34.5137 | -17.6016
10.2674 | -34.5141 | -17.6111
10.2735 | -34.5145 | -17.6207
10.2796 | -34.515 | -17.6301
10.2857 | -34.5154 | -17.6398
10.2918 | -34.5158 | -17.6493
10.2979 | -34.5163 | -17.6588
10.304 | -34.5167 | -17.668
10.3101 | -34.5171 | -17.6776
10.3162 | -34.5175 | -17.6871
10.3223 | -34.518 | -17.6964
10.3284 | -34.5184 | -17.7058
10.3345 | -34.5188 | -17.7152
10.3406 | -34.5192 | -17.7246
10.3467 | -34.5197 | -17.7339
10.3528 | -34.5201 | -17.7433
10.3589 | -34.5205 | -17.7528
10.365 | -34.5209 | -17.7624
10.3711 | -34.5213 | -17.7719
10.3772 | -34.5218 | -17.7813
10.3833 | -34.5222 | -17.7907
10.3895 | -34.5227 | -17.8002
10.3956 | -34.5231 | -17.8096
10.4017 | -34.5235 | -17.8189
10.4078 | -34.5239 | -17.8284
10.4139 | -34.5244 | -17.8378
10.42 | -34.5248 | -17.8474
10.4261 | -34.5252 | -17.8568
10.4322 | -34.5256 | -17.8662
10.4383 | -34.526 | -17.8757
10.4444 | -34.5264 | -17.8849
10.4505 | -34.5269 | -17.8944
10.4566 | -34.5273 | -17.9037
10.4627 | -34.5277 | -17.913
10.4688 | -34.5282 | -17.9223
10.4749 | -34.5286 | -17.9318
10.481 | -34.529 | -17.9412
10.4871 | -34.5294 | -17.9505
10.4932 | -34.5299 | -17.9598
10.4993 | -34.5303 | -17.9693
10.5054 | -34.5307 | -17.9787
10.5115 | -34.5311 | -17.9881
10.5176 | -34.5316 | -17.9973
10.5237 | -34.5321 | -18.0069
10.5298 | -34.5325 | -18.0161
10.536 | -34.5329 | -18.0256
10.5421 | -34.5333 | -18.0348
10.5482 | -34.5337 | -18.0442
10.5543 | -34.5342 | -18.0533
10.5604 | -34.5346 | -18.0626
10.5665 | -34.535 | -18.0718
10.5726 | -34.5355 | -18.0811
10.5787 | -34.5359 | -18.0905
10.5848 | -34.5363 | -18.0995
10.5909 | -34.5367 | -18.1089
10.597 | -34.5372 | -18.1183
10.6031 | -34.5376 | -18.1278
10.6092 | -34.538 | -18.1371
10.6153 | -34.5384 | -18.1465
10.6214 | -34.5389 | -18.1555
10.6275 | -34.5393 | -18.1648
10.6336 | -34.5397 | -18.1743
10.6397 | -34.5401 | -18.1837
10.6458 | -34.5406 | -18.1929
10.6519 | -34.541 | -18.2023
10.658 | -34.5414 | -18.2116
10.6641 | -34.5419 | -18.2211
10.6702 | -34.5423 | -18.2304
10.6764 | -34.5427 | -18.2398
10.6825 | -34.5432 | -18.249
10.6886 | -34.5436 | -18.2582
10.6947 | -34.544 | -18.2674
10.7008 | -34.5445 | -18.2768
10.7069 | -34.5449 | -18.2862
10.713 | -34.5453 | -18.2955
10.7191 | -34.5458 | -18.3049
10.7252 | -34.5462 | -18.3139
10.7313 | -34.5466 | -18.3231
10.7374 | -34.547 | -18.3325
10.7435 | -34.5475 | -18.3417
10.7496 | -34.5479 | -18.3508
10.7557 | -34.5483 | -18.3602
10.7618 | -34.5488 | -18.3691
10.7679 | -34.5492 | -18.3785
10.774 | -34.5496 | -18.3878
10.7801 | -34.55 | -18.3972
10.7862 | -34.5505 | -18.4064
10.7923 | -34.5509 | -18.4158
10.7984 | -34.5514 | -18.4249
10.8045 | -34.5518 | -18.4341
10.8106 | -34.5522 | -18.4435
10.8168 | -34.5526 | -18.4526
10.8229 | -34.5531 | -18.4619
10.829 | -34.5535 | -18.4711
10.8351 | -34.5539 | -18.48
10.8412 | -34.5544 | -18.4893
10.8473 | -34.5548 | -18.4984
10.8534 | -34.5552 | -18.5076
10.8595 | -34.5557 | -18.5169
10.8656 | -34.5561 | -18.5258
10.8717 | -34.5565 | -18.5351
10.8778 | -34.557 | -18.5444
10.8839 | -34.5574 | -18.5536
10.89 | -34.5578 | -18.5628
10.8961 | -34.5583 | -18.572
10.9022 | -34.5587 | -18.5811
10.9083 | -34.5592 | -18.5902
10.9144 | -34.5596 | -18.5994
10.9205 | -34.56 | -18.6086
10.9266 | -34.5605 | -18.6178
10.9327 | -34.5609 | -18.6269
10.9388 | -34.5613 | -18.636
10.9449 | -34.5618 | -18.6451
10.951 | -34.5622 | -18.6544
10.9571 | -34.5626 | -18.6636
10.9633 | -34.5631 | -18.6729
10.9694 | -34.5635 | -18.6823
10.9755 | -34.5639 | -18.6915
10.9816 | -34.5644 | -18.7008
10.9877 | -34.5648 | -18.7099
10.9938 | -34.5653 | -18.7193
10.9999 | -34.5657 | -18.7284
11.006 | -34.5661 | -18.7374
11.0121 | -34.5666 | -18.7466
11.0182 | -34.567 | -18.7557
11.0243 | -34.5674 | -18.7649
11.0304 | -34.5679 | -18.7741
11.0365 | -34.5683 | -18.7832
11.0426 | -34.5688 | -18.7923
11.0487 | -34.5692 | -18.8016
11.0548 | -34.5697 | -18.8107
11.0609 | -34.5701 | -18.8201
11.067 | -34.5706 | -18.8292
11.0731 | -34.571 | -18.8383
11.0792 | -34.5715 | -18.8474
11.0853 | -34.5719 | -18.8566
11.0914 | -34.5724 | -18.866
11.0975 | -34.5728 | -18.8751
11.1037 | -34.5732 | -18.8842
11.1098 | -34.5736 | -18.8935
11.1159 | -34.5741 | -18.9023
11.122 | -34.5745 | -18.9115
11.1281 | -34.575 | -18.9208
11.1342 | -34.5754 | -18.9299
11.1403 | -34.5759 | -18.9391
11.1464 | -34.5763 | -18.9483
11.1525 | -34.5767 | -18.9572
11.1586 | -34.5772 | -18.9664
11.1647 | -34.5776 | -18.9754
11.1708 | -34.578 | -18.9845
11.1769 | -34.5785 | -18.9937
11.183 | -34.5789 | -19.0027
11.1891 | -34.5794 | -19.0119
11.1952 | -34.5798 | -19.021
11.2013 | -34.5803 | -19.0301
11.2074 | -34.5807 | -19.0392
11.2135 | -34.5812 | -19.0485
11.2196 | -34.5816 | -19.0574
11.2257 | -34.5821 | -19.0666
11.2318 | -34.5825 | -19.0756
11.2379 | -34.583 | -19.0848
11.244 | -34.5835 | -19.094
11.2502 | -34.5839 | -19.1031
11.2563 | -34.5844 | -19.1121
11.2624 | -34.5848 | -19.1211
11.2685 | -34.5853 | -19.1302
11.2746 | -34.5857 | -19.1392
11.2807 | -34.5861 | -19.1485
11.2868 | -34.5866 | -19.1577
11.2929 | -34.587 | -19.1666
11.299 | -34.5875 | -19.1759
11.3051 | -34.5879 | -19.185
11.3112 | -34.5884 | -19.1943
11.3173 | -34.5889 | -19.2033
11.3234 | -34.5893 | -19.2122
11.3295 | -34.5898 | -19.2214
11.3356 | -34.5902 | -19.2307
11.3417 | -34.5907 | -19.2397
11.3478 | -34.5911 | -19.2488
11.3539 | -34.5915 | -19.258
11.36 | -34.592 | -19.267
11.3661 | -34.5924 | -19.276
11.3722 | -34.5929 | -19.285
11.3783 | -34.5933 | -19.2942
11.3844 | -34.5938 | -19.3033
11.3906 | -34.5942 | -19.3123
11.3967 | -34.5947 | -19.3212
11.4028 | -34.5951 | -19.3305
11.4089 | -34.5956 | -19.3394
11.415 | -34.596 | -19.3486
11.4211 | -34.5964 | -19.3578
11.4272 | -34.5969 | -19.3669
11.4333 | -34.5973 | -19.3758
11.4394 | -34.5978 | -19.3849
11.4455 | -34.5982 | -19.3939
11.4516 | -34.5986 | -19.4029
11.4577 | -34.5991 | -19.412
11.4638 | -34.5995 | -19.4208
11.4699 | -34.6 | -19.43
11.476 | -34.6004 | -19.4392
11.4821 | -34.6009 | -19.4482
11.4882 | -34.6013 | -19.4574
11.4943 | -34.6018 | -19.4665
11.5004 | -34.6023 | -19.4754
11.5065 | -34.6027 | -19.4844
11.5126 | -34.6032 | -19.4934
11.5187 | -34.6036 | -19.5025
11.5248 | -34.6041 | -19.5118
11.5309 | -34.6045 | -19.5207
11.5371 | -34.605 | -19.5297
11.5432 | -34.6054 | -19.5388
11.5493 | -34.6059 | -19.548
11.5554 | -34.6063 | -19.5571
11.5615 | -34.6068 | -19.5663
11.5676 | -34.6072 | -19.5753
11.5737 | -34.6077 | -19.5841
11.5798 | -34.6081 | -19.593
11.5859 | -34.6086 | -19.602
11.592 | -34.609 | -19.611
11.5981 | -34.6095 | -19.6199
11.6042 | -34.61 | -19.629
11.6103 | -34.6104 | -19.6379
11.6164 | -34.6108 | -19.6468
11.6225 | -34.6113 | -19.6558
11.6286 | -34.6117 | -19.6649
11.6347 | -34.6122 | -19.674
11.6408 | -34.6127 | -19.6829
11.6469 | -34.6131 | -19.6919
11.653 | -34.6136 | -19.7009
11.6591 | -34.614 | -19.7099
11.6652 | -34.6145 | -19.7189
11.6713 | -34.615 | -19.7278
11.6775 | -34.6154 | -19.7366
11.6836 | -34.6159 | -19.7455
11.6897 | -34.6164 | -19.7545
11.6958 | -34.6168 | -19.7636
11.7019 | -34.6172 | -19.7727
11.708 | -34.6177 | -19.7817
11.7141 | -34.6182 | -19.7907
11.7202 | -34.6186 | -19.7998
11.7263 | -34.619 | -19.8088
11.7324 | -34.6195 | -19.818
11.7385 | -34.6199 | -19.827
11.7446 | -34.6204 | -19.8359
11.7507 | -34.6208 | -19.8446
11.7568 | -34.6213 | -19.8536
11.7629 | -34.6217 | -19.8627
11.769 | -34.6222 | -19.8716
11.7751 | -34.6226 | -19.8805
11.7812 | -34.6231 | -19.8894
11.7873 | -34.6236 | -19.8984
11.7934 | -34.624 | -19.9074
11.7995 | -34.6245 | -19.9162
11.8056 | -34.6249 | -19.9252
11.8117 | -34.6254 | -19.9341
11.8178 | -34.6259 | -19.9431
11.824 | -34.6263 | -19.952
11.8301 | -34.6268 | -19.9609
11.8362 | -34.6272 | -19.9697
11.8423 | -34.6277 | -19.9787
11.8484 | -34.6281 | -19.9877
11.8545 | -34.6286 | -19.9964
11.8606 | -34.629 | -20.0053
11.8667 | -34.6295 | -20.0144
11.8728 | -34.63 | -20.0233
11.8789 | -34.6304 | -20.0322
11.885 | -34.6309 | -20.0412
11.8911 | -34.6313 | -20.0502
11.8972 | -34.6318 | -20.0591
11.9033 | -34.6322 | -20.068
11.9094 | -34.6327 | -20.0769
11.9155 | -34.6331 | -20.0858
11.9216 | -34.6336 | -20.0946
11.9277 | -34.6341 | -20.1034
11.9338 | -34.6345 | -20.1125
11.9399 | -34.635 | -20.1215
11.946 | -34.6354 | -20.1303
11.9521 | -34.6359 | -20.1392
11.9582 | -34.6363 | -20.148
11.9644 | -34.6368 | -20.1568
11.9705 | -34.6373 | -20.1658
11.9766 | -34.6377 | -20.1747
11.9827 | -34.6382 | -20.1835
11.9888 | -34.6387 | -20.1924
11.9949 | -34.6392 | -20.201
12.001 | -34.6396 | -20.2099
12.0071 | -34.6401 | -20.2188
12.0132 | -34.6405 | -20.2277
12.0193 | -34.6409 | -20.2367
12.0254 | -34.6414 | -20.2457
12.0315 | -34.6419 | -20.2543
12.0376 | -34.6423 | -20.2632
12.0437 | -34.6427 | -20.2721
12.0498 | -34.6432 | -20.281
12.0559 | -34.6437 | -20.2897
12.062 | -34.6441 | -20.2985
12.0681 | -34.6446 | -20.3075
12.0742 | -34.6451 | -20.3162
12.0803 | -34.6455 | -20.3251
12.0864 | -34.646 | -20.334
12.0925 | -34.6464 | -20.3429
12.0986 | -34.6469 | -20.3517
12.1047 | -34.6473 | -20.3606
12.1109 | -34.6478 | -20.3694
12.117 | -34.6483 | -20.3782
12.1231 | -34.6487 | -20.387
12.1292 | -34.6492 | -20.3958
12.1353 | -34.6496 | -20.4046
12.1414 | -34.6501 | -20.4134
12.1475 | -34.6506 | -20.4224
12.1536 | -34.651 | -20.4313
12.1597 | -34.6515 | -20.4402
12.1658 | -34.652 | -20.4489
12.1719 | -34.6524 | -20.4576
12.178 | -34.6529 | -20.4665
12.1841 | -34.6533 | -20.4755
12.1902 | -34.6538 | -20.4844
12.1963 | -34.6543 | -20.4933
12.2024 | -34.6547 | -20.502
12.2085 | -34.6552 | -20.5109
12.2146 | -34.6557 | -20.5197
12.2207 | -34.6561 | -20.5286
12.2268 | -34.6566 | -20.5375
12.2329 | -34.657 | -20.5463
12.239 | -34.6575 | -20.5549
12.2451 | -34.658 | -20.5637
12.2513 | -34.6584 | -20.5727
12.2574 | -34.6589 | -20.5816
12.2635 | -34.6594 | -20.5904
12.2696 | -34.6598 | -20.5992
12.2757 | -34.6603 | -20.6077
12.2818 | -34.6608 | -20.6165
12.2879 | -34.6612 | -20.6254
12.294 | -34.6617 | -20.6341
12.3001 | -34.6621 | -20.6429
12.3062 | -34.6626 | -20.6517
12.3123 | -34.6631 | -20.6603
12.3184 | -34.6635 | -20.669
12.3245 | -34.664 | -20.6777
12.3306 | -34.6644 | -20.6865
12.3367 | -34.6649 | -20.6953
12.3428 | -34.6653 | -20.7041
12.3489 | -34.6658 | -20.7127
12.355 | -34.6663 | -20.7216
12.3611 | -34.6667 | -20.7305
12.3672 | -34.6672 | -20.7392
12.3733 | -34.6677 | -20.748
12.3794 | -34.6681 | -20.7565
12.3855 | -34.6686 | -20.7652
12.3916 | -34.669 | -20.7739
12.3978 | -34.6695 | -20.7827
12.4039 | -34.67 | -20.7915
12.41 | -34.6705 | -20.8004
12.4161 | -34.6709 | -20.8089
12.4222 | -34.6714 | -20.8177
12.4283 | -34.6719 | -20.8264
12.4344 | -34.6723 | -20.8352
12.4405 | -34.6728 | -20.844
12.4466 | -34.6732 | -20.8527
12.4527 | -34.6737 | -20.8614
12.4588 | -34.6742 | -20.8701
12.4649 | -34.6746 | -20.879
12.471 | -34.6751 | -20.8879
12.4771 | -34.6755 | -20.8966
12.4832 | -34.676 | -20.9053
12.4893 | -34.6765 | -20.9138
12.4954 | -34.6769 | -20.9224
12.5015 | -34.6774 | -20.9311
12.5076 | -34.6778 | -20.94
12.5137 | -34.6783 | -20.9487
12.5198 | -34.6788 | -20.9573
12.5259 | -34.6792 | -20.9662
12.532 | -34.6797 | -20.9748
12.5382 | -34.6801 | -20.9836
12.5443 | -34.6806 | -20.9923
12.5504 | -34.681 | -21.0011
12.5565 | -34.6815 | -21.0097
12.5626 | -34.682 | -21.0186
12.5687 | -34.6825 | -21.0273
12.5748 | -34.6829 | -21.0361
12.5809 | -34.6833 | -21.0448
12.587 | -34.6838 | -21.0536
12.5931 | -34.6843 | -21.0622
12.5992 | -34.6848 | -21.071
12.6053 | -34.6853 | -21.0796
12.6114 | -34.6857 | -21.0883
12.6175 | -34.6862 | -21.0971
12.6236 | -34.6866 | -21.1058
12.6297 | -34.6871 | -21.1145
12.6358 | -34.6875 | -21.1231
12.6419 | -34.688 | -21.1318
12.648 | -34.6885 | -21.1406
12.6541 | -34.6889 | -21.1493
12.6602 | -34.6894 | -21.1577
12.6663 | -34.6898 | -21.1665
12.6724 | -34.6903 | -21.1751
12.6785 | -34.6907 | -21.1838
12.6847 | -34.6912 | -21.1924
12.6908 | -34.6917 | -21.2011
12.6969 | -34.6921 | -21.2097
12.703 | -34.6926 | -21.2185
12.7091 | -34.693 | -21.2271
12.7152 | -34.6935 | -21.2357
12.7213 | -34.694 | -21.2444
12.7274 | -34.6944 | -21.2531
12.7335 | -34.6949 | -21.2615
12.7396 | -34.6954 | -21.27
12.7457 | -34.6958 | -21.2788
12.7518 | -34.6963 | -21.2875
12.7579 | -34.6967 | -21.2963
12.764 | -34.6972 | -21.305
12.7701 | -34.6976 | -21.3137
12.7762 | -34.6981 | -21.3222
12.7823 | -34.6986 | -21.3309
12.7884 | -34.6991 | -21.3396
12.7945 | -34.6995 | -21.3485
12.8006 | -34.7 | -21.357
12.8067 | -34.7005 | -21.3656
12.8128 | -34.7009 | -21.3741
12.8189 | -34.7014 | -21.3828
12.8251 | -34.7018 | -21.3915
12.8312 | -34.7023 | -21.4003
12.8373 | -34.7028 | -21.4088
12.8434 | -34.7032 | -21.4175
12.8495 | -34.7037 | -21.4261
12.8556 | -34.7041 | -21.4347
12.8617 | -34.7046 | -21.4435
12.8678 | -34.705 | -21.4523
12.8739 | -34.7055 | -21.4608
12.88 | -34.706 | -21.4694
12.8861 | -34.7064 | -21.478
12.8922 | -34.7069 | -21.4867
12.8983 | -34.7073 | -21.4953
12.9044 | -34.7078 | -21.5039
12.9105 | -34.7083 | -21.5124
12.9166 | -34.7087 | -21.521
12.9227 | -34.7092 | -21.5297
12.9288 | -34.7097 | -21.5383
12.9349 | -34.7101 | -21.547
12.941 | -34.7106 | -21.5556
12.9471 | -34.711 | -21.5642
12.9532 | -34.7115 | -21.5728
12.9593 | -34.712 | -21.5815
12.9654 | -34.7125 | -21.5903
12.9716 | -34.7129 | -21.599
12.9777 | -34.7134 | -21.6076
12.9838 | -34.7139 | -21.6162
12.9899 | -34.7143 | -21.6248
12.996 | -34.7147 | -21.6335
13.0021 | -34.7152 | -21.6421
13.0082 | -34.7157 | -21.6507
13.0143 | -34.7162 | -21.6591
13.0204 | -34.7166 | -21.6676
13.0265 | -34.7171 | -21.6761
13.0326 | -34.7175 | -21.6848
13.0387 | -34.718 | -21.6936
13.0448 | -34.7185 | -21.7021
13.0509 | -34.7189 | -21.7107
13.057 | -34.7194 | -21.7192
13.0631 | -34.7199 | -21.7277
13.0692 | -34.7204 | -21.7363
13.0753 | -34.7208 | -21.7449
13.0814 | -34.7213 | -21.7537
13.0875 | -34.7218 | -21.7621
13.0936 | -34.7222 | -21.7708
13.0997 | -34.7227 | -21.7793
13.1058 | -34.7231 | -21.7879
13.112 | -34.7236 | -21.7965
13.1181 | -34.7241 | -21.8051
13.1242 | -34.7245 | -21.8137
13.1303 | -34.725 | -21.8223
13.1364 | -34.7254 | -21.8309
13.1425 | -34.7259 | -21.8395
13.1486 | -34.7263 | -21.8482
13.1547 | -34.7268 | -21.8567
13.1608 | -34.7273 | -21.8652
13.1669 | -34.7277 | -21.8737
13.173 | -34.7282 | -21.8823
13.1791 | -34.7287 | -21.8908
13.1852 | -34.7291 | -21.8994
13.1913 | -34.7296 | -21.9079
13.1974 | -34.73 | -21.9165
13.2035 | -34.7305 | -21.9251
13.2096 | -34.731 | -21.9336
13.2157 | -34.7314 | -21.9421
13.2218 | -34.7319 | -21.9507
13.2279 | -34.7324 | -21.9592
13.234 | -34.7329 | -21.9677
13.2401 | -34.7333 | -21.9762
13.2462 | -34.7338 | -21.9847
13.2524 | -34.7342 | -21.9933
13.2585 | -34.7347 | -22.0019
13.2646 | -34.7351 | -22.0104
13.2707 | -34.7356 | -22.019
13.2768 | -34.7361 | -22.0277
13.2829 | -34.7365 | -22.0363
13.289 | -34.737 | -22.0448
13.2951 | -34.7375 | -22.0532
13.3012 | -34.7379 | -22.0618
13.3073 | -34.7384 | -22.0703
13.3134 | -34.7389 | -22.079
13.3195 | -34.7393 | -22.0875
13.3256 | -34.7398 | -22.0959
13.3317 | -34.7403 | -22.1043
13.3378 | -34.7407 | -22.1129
13.3439 | -34.7412 | -22.1216
13.35 | -34.7417 | -22.1303
13.3561 | -34.7421 | -22.1387
13.3622 | -34.7426 | -22.1471
13.3683 | -34.743 | -22.1555
13.3744 | -34.7435 | -22.164
13.3805 | -34.744 | -22.1726
13.3866 | -34.7444 | -22.1811
13.3927 | -34.7449 | -22.1896
13.3989 | -34.7453 | -22.1981
13.405 | -34.7458 | -22.2067
13.4111 | -34.7462 | -22.2152
13.4172 | -34.7467 | -22.2238
13.4233 | -34.7472 | -22.2324
13.4294 | -34.7476 | -22.2412
13.4355 | -34.7481 | -22.2494
13.4416 | -34.7486 | -22.2579
13.4477 | -34.749 | -22.2664
13.4538 | -34.7495 | -22.275
13.4599 | -34.75 | -22.2836
13.466 | -34.7505 | -22.2922
13.4721 | -34.7509 | -22.3007
13.4782 | -34.7514 | -22.3092
13.4843 | -34.7519 | -22.3178
13.4904 | -34.7523 | -22.3263
13.4965 | -34.7528 | -22.335
13.5026 | -34.7532 | -22.3436
13.5087 | -34.7537 | -22.352
13.5148 | -34.7542 | -22.3607
13.5209 | -34.7546 | -22.3693
13.527 | -34.7551 | -22.3778
13.5331 | -34.7556 | -22.3863
13.5393 | -34.756 | -22.3947
13.5454 | -34.7565 | -22.4032
13.5515 | -34.757 | -22.4119
13.5576 | -34.7575 | -22.4204
13.5637 | -34.7579 | -22.429
13.5698 | -34.7584 | -22.4375
13.5759 | -34.7589 | -22.4458
13.582 | -34.7593 | -22.4542
13.5881 | -34.7598 | -22.4627
13.5942 | -34.7603 | -22.4714
13.6003 | -34.7607 | -22.4798
13.6064 | -34.7612 | -22.4884
13.6125 | -34.7617 | -22.4969
13.6186 | -34.7621 | -22.5054
13.6247 | -34.7626 | -22.514
13.6308 | -34.7631 | -22.5226
13.6369 | -34.7635 | -22.5311
13.643 | -34.764 | -22.5398
13.6491 | -34.7644 | -22.5482
13.6552 | -34.7649 | -22.5567
13.6613 | -34.7654 | -22.5652
13.6674 | -34.7659 | -22.5737
13.6735 | -34.7663 | -22.5822
13.6796 | -34.7668 | -22.5907
13.6858 | -34.7672 | -22.5992
13.6919 | -34.7677 | -22.6077
13.698 | -34.7682 | -22.6162
13.7041 | -34.7687 | -22.6249
13.7102 | -34.7691 | -22.6333
13.7163 | -34.7696 | -22.6418
13.7224 | -34.7701 | -22.6504
13.7285 | -34.7706 | -22.6589
13.7346 | -34.771 | -22.6674
13.7407 | -34.7715 | -22.6759
13.7468 | -34.7719 | -22.6844
13.7529 | -34.7724 | -22.6927
13.759 | -34.7729 | -22.7014
13.7651 | -34.7733 | -22.7099
13.7712 | -34.7738 | -22.7184
13.7773 | -34.7743 | -22.727
13.7834 | -34.7747 | -22.7355
13.7895 | -34.7752 | -22.744
13.7956 | -34.7757 | -22.7525
13.8017 | -34.7761 | -22.7609
13.8078 | -34.7766 | -22.7695
13.8139 | -34.7771 | -22.7779
13.82 | -34.7775 | -22.7863
13.8262 | -34.778 | -22.7948
13.8323 | -34.7785 | -22.8034
13.8384 | -34.7789 | -22.8119
13.8445 | -34.7794 | -22.8204
13.8506 | -34.7799 | -22.8289
13.8567 | -34.7804 | -22.8375
13.8628 | -34.7808 | -22.8461
13.8689 | -34.7813 | -22.8547
13.875 | -34.7818 | -22.8632
13.8811 | -34.7822 | -22.8717
13.8872 | -34.7826 | -22.8803
13.8933 | -34.7831 | -22.8887
13.8994 | -34.7836 | -22.8972
13.9055 | -34.7841 | -22.9058
13.9116 | -34.7845 | -22.9143
13.9177 | -34.785 | -22.9229
13.9238 | -34.7855 | -22.9314
13.9299 | -34.786 | -22.9397
13.936 | -34.7864 | -22.9482
13.9421 | -34.7869 | -22.9566
13.9482 | -34.7873 | -22.965
13.9543 | -34.7878 | -22.9735
13.9604 | -34.7882 | -22.982
13.9665 | -34.7887 | -22.9904
13.9727 | -34.7892 | -22.9991
13.9788 | -34.7897 | -23.0075
13.9849 | -34.7901 | -23.016
13.991 | -34.7906 | -23.0245
13.9971 | -34.7911 | -23.0328
14.0032 | -34.7915 | -23.0412
14.0093 | -34.792 | -23.0497
14.0154 | -34.7925 | -23.0581
14.0215 | -34.7929 | -23.0665
14.0276 | -34.7934 | -23.075
14.0337 | -34.7939 | -23.0834
14.0398 | -34.7944 | -23.0918
14.0459 | -34.7949 | -23.1003
14.052 | -34.7953 | -23.1088
14.0581 | -34.7958 | -23.1173
14.0642 | -34.7962 | -23.1259
14.0703 | -34.7967 | -23.1342
14.0764 | -34.7972 | -23.1426
14.0825 | -34.7977 | -23.1511
14.0886 | -34.7981 | -23.1596
14.0947 | -34.7986 | -23.1681
14.1008 | -34.7991 | -23.1765
14.1069 | -34.7996 | -23.1848
14.1131 | -34.8 | -23.1933
14.1192 | -34.8005 | -23.2016
14.1253 | -34.801 | -23.21
14.1314 | -34.8015 | -23.2184
14.1375 | -34.8019 | -23.2268
14.1436 | -34.8024 | -23.2352
14.1497 | -34.8029 | -23.2436
14.1558 | -34.8034 | -23.252
14.1619 | -34.8038 | -23.2604
14.168 | -34.8043 | -23.2689
14.1741 | -34.8048 | -23.2772
14.1802 | -34.8053 | -23.2856
14.1863 | -34.8057 | -23.2941
14.1924 | -34.8062 | -23.3026
14.1985 | -34.8067 | -23.311
14.2046 | -34.8072 | -23.3195
14.2107 | -34.8077 | -23.3278
14.2168 | -34.8081 | -23.3362
14.2229 | -34.8086 | -23.3447
14.229 | -34.8091 | -23.3532
14.2351 | -34.8095 | -23.3616
14.2412 | -34.81 | -23.3701
14.2473 | -34.8105 | -23.3784
14.2534 | -34.811 | -23.3868
14.2596 | -34.8114 | -23.3951
14.2657 | -34.8119 | -23.4036
14.2718 | -34.8123 | -23.4121
14.2779 | -34.8128 | -23.4206
14.284 | -34.8133 | -23.429
14.2901 | -34.8138 | -23.4374
14.2962 | -34.8143 | -23.4459
14.3023 | -34.8148 | -23.4543
14.3084 | -34.8152 | -23.4627
14.3145 | -34.8157 | -23.471
14.3206 | -34.8162 | -23.4796
14.3267 | -34.8167 | -23.4881
14.3328 | -34.8172 | -23.4965
14.3389 | -34.8176 | -23.5048
14.345 | -34.8181 | -23.5133
14.3511 | -34.8186 | -23.5216
14.3572 | -34.8191 | -23.5301
14.3633 | -34.8195 | -23.5386
14.3694 | -34.82 | -23.5469
14.3755 | -34.8205 | -23.5553
14.3816 | -34.821 | -23.5638
14.3877 | -34.8215 | -23.5722
14.3938 | -34.822 | -23.5808
14.4 | -34.8224 | -23.5892
14.4061 | -34.8229 | -23.5976
14.4122 | -34.8234 | -23.606
14.4183 | -34.8238 | -23.6144
14.4244 | -34.8243 | -23.6228
14.4305 | -34.8248 | -23.6312
14.4366 | -34.8253 | -23.6396
14.4427 | -34.8258 | -23.6482
14.4488 | -34.8263 | -23.6566
14.4549 | -34.8267 | -23.665
14.461 | -34.8272 | -23.6732
14.4671 | -34.8277 | -23.6818
14.4732 | -34.8282 | -23.6903
14.4793 | -34.8287 | -23.6988
14.4854 | -34.8291 | -23.7071
14.4915 | -34.8296 | -23.7154
14.4976 | -34.8301 | -23.7237
14.5037 | -34.8306 | -23.7321
14.5098 | -34.8311 | -23.7404
14.5159 | -34.8315 | -23.7488
14.522 | -34.832 | -23.7571
14.5281 | -34.8325 | -23.7655
14.5342 | -34.833 | -23.7739
14.5403 | -34.8335 | -23.7822
14.5465 | -34.8339 | -23.7906
14.5526 | -34.8344 | -23.7991
14.5587 | -34.8349 | -23.8075
14.5648 | -34.8354 | -23.8158
14.5709 | -34.8359 | -23.8241
14.577 | -34.8363 | -23.8324
14.5831 | -34.8368 | -23.8408
14.5892 | -34.8373 | -23.8491
14.5953 | -34.8378 | -23.8574
14.6014 | -34.8383 | -23.8657
14.6075 | -34.8387 | -23.874
14.6136 | -34.8392 | -23.8824
14.6197 | -34.8397 | -23.8908
14.6258 | -34.8401 | -23.8991
14.6319 | -34.8406 | -23.9074
14.638 | -34.8411 | -23.9157
14.6441 | -34.8415 | -23.9242
14.6502 | -34.842 | -23.9326
14.6563 | -34.8425 | -23.9409
14.6624 | -34.843 | -23.9493
14.6685 | -34.8435 | -23.9574
14.6746 | -34.844 | -23.9657
14.6807 | -34.8444 | -23.974
14.6869 | -34.8449 | -23.9823
14.693 | -34.8454 | -23.9906
14.6991 | -34.8459 | -23.9989
14.7052 | -34.8464 | -24.0071
14.7113 | -34.8469 | -24.0155
14.7174 | -34.8474 | -24.0238
14.7235 | -34.8479 | -24.0321
14.7296 | -34.8483 | -24.0403
14.7357 | -34.8489 | -24.0485
14.7418 | -34.8493 | -24.0569
14.7479 | -34.8498 | -24.0651
14.754 | -34.8503 | -24.0735
14.7601 | -34.8508 | -24.0817
14.7662 | -34.8512 | -24.0899
14.7723 | -34.8517 | -24.0981
14.7784 | -34.8522 | -24.1064
14.7845 | -34.8527 | -24.1147
14.7906 | -34.8532 | -24.1231
14.7967 | -34.8537 | -24.1313
14.8028 | -34.8541 | -24.1396
14.8089 | -34.8547 | -24.1478
14.815 | -34.8551 | -24.1561
14.8211 | -34.8556 | -24.1644
14.8272 | -34.8561 | -24.1728
14.8334 | -34.8566 | -24.181
14.8395 | -34.8571 | -24.1893
14.8456 | -34.8576 | -24.1975
14.8517 | -34.8581 | -24.2058
14.8578 | -34.8585 | -24.2142
14.8639 | -34.859 | -24.2225
14.87 | -34.8595 | -24.2309
14.8761 | -34.86 | -24.239
14.8822 | -34.8605 | -24.2474
14.8883 | -34.861 | -24.2557
14.8944 | -34.8615 | -24.2639
14.9005 | -34.862 | -24.2722
14.9066 | -34.8625 | -24.2806
14.9127 | -34.863 | -24.2888
14.9188 | -34.8634 | -24.2971
14.9249 | -34.8639 | -24.3054
14.931 | -34.8644 | -24.3137
14.9371 | -34.8649 | -24.3218
14.9432 | -34.8654 | -24.3302
14.9493 | -34.8659 | -24.3384
14.9554 | -34.8664 | -24.3466
14.9615 | -34.8668 | -24.3549
14.9676 | -34.8673 | -24.3633
14.9738 | -34.8678 | -24.3714
14.9799 | -34.8683 | -24.3799
14.986 | -34.8687 | -24.388
14.9921 | -34.8692 | -24.3963
14.9982 | -34.8697 | -24.4046
15.0043 | -34.8702 | -24.4128
15.0104 | -34.8707 | -24.4211
15.0165 | -34.8712 | -24.4295
15.0226 | -34.8717 | -24.4375
15.0287 | -34.8722 | -24.4458
15.0348 | -34.8727 | -24.4541
15.0409 | -34.8732 | -24.4625
15.047 | -34.8737 | -24.4709
15.0531 | -34.8742 | -24.4789
15.0592 | -34.8747 | -24.4872
15.0653 | -34.8752 | -24.4956
15.0714 | -34.8757 | -24.5039
15.0775 | -34.8761 | -24.512
15.0836 | -34.8767 | -24.5203
15.0897 | -34.8772 | -24.5283
15.0958 | -34.8777 | -24.5365
15.1019 | -34.8782 | -24.5448
15.108 | -34.8787 | -24.553
15.1141 | -34.8791 | -24.5612
15.1203 | -34.8796 | -24.5694
15.1264 | -34.8801 | -24.5776
15.1325 | -34.8806 | -24.5859
15.1386 | -34.8811 | -24.5941
15.1447 | -34.8816 | -24.6023
15.1508 | -34.8821 | -24.6106
15.1569 | -34.8826 | -24.6188
15.163 | -34.8831 | -24.6268
15.1691 | -34.8836 | -24.6349
15.1752 | -34.8841 | -24.6431
15.1813 | -34.8846 | -24.6512
15.1874 | -34.8851 | -24.6596
15.1935 | -34.8856 | -24.6675
15.1996 | -34.886 | -24.6756
15.2057 | -34.8865 | -24.6838
15.2118 | -34.887 | -24.692
15.2179 | -34.8875 | -24.7002
15.224 | -34.888 | -24.7083
15.2301 | -34.8885 | -24.7162
15.2362 | -34.889 | -24.7243
15.2423 | -34.8895 | -24.7326
15.2484 | -34.89 | -24.7408
15.2545 | -34.8905 | -24.749
15.2607 | -34.891 | -24.7572
15.2668 | -34.8915 | -24.7651
15.2729 | -34.892 | -24.7732
15.279 | -34.8925 | -24.7817
15.2851 | -34.893 | -24.7898
15.2912 | -34.8935 | -24.7981
15.2973 | -34.894 | -24.8063
15.3034 | -34.8945 | -24.8143
15.3095 | -34.895 | -24.8225
15.3156 | -34.8955 | -24.8306
15.3217 | -34.896 | -24.8387
15.3278 | -34.8965 | -24.847
15.3339 | -34.897 | -24.855
15.34 | -34.8975 | -24.8631
15.3461 | -34.898 | -24.8712
15.3522 | -34.8985 | -24.8793
15.3583 | -34.899 | -24.8876
15.3644 | -34.8995 | -24.8957
15.3705 | -34.9 | -24.9035
15.3766 | -34.9005 | -24.9118
15.3827 | -34.901 | -24.9198
15.3888 | -34.9015 | -24.928
15.3949 | -34.902 | -24.9361
15.401 | -34.9025 | -24.9444
15.4072 | -34.903 | -24.9522
15.4133 | -34.9035 | -24.9603
15.4194 | -34.904 | -24.9686
15.4255 | -34.9045 | -24.9768
15.4316 | -34.905 | -24.985
15.4377 | -34.9055 | -24.9933
15.4438 | -34.906 | -25.0013
15.4499 | -34.9065 | -25.0094
15.456 | -34.907 | -25.0174
15.4621 | -34.9075 | -25.0255
15.4682 | -34.908 | -25.0337
15.4743 | -34.9085 | -25.0416
15.4804 | -34.909 | -25.0498
15.4865 | -34.9095 | -25.0579
15.4926 | -34.91 | -25.0661
15.4987 | -34.9105 | -25.0742
15.5048 | -34.911 | -25.0824
15.5109 | -34.9115 | -25.0904
15.517 | -34.912 | -25.0985
15.5231 | -34.9125 | -25.1065
15.5292 | -34.913 | -25.1146
15.5353 | -34.9135 | -25.1229
15.5414 | -34.914 | -25.1309
15.5476 | -34.9145 | -25.1387
15.5537 | -34.915 | -25.1468
15.5598 | -34.9155 | -25.155
15.5659 | -34.916 | -25.163
15.572 | -34.9166 | -25.171
15.5781 | -34.9171 | -25.179
15.5842 | -34.9176 | -25.1869
15.5903 | -34.9181 | -25.1949
15.5964 | -34.9186 | -25.2028
15.6025 | -34.9191 | -25.2108
15.6086 | -34.9196 | -25.2189
15.6147 | -34.9201 | -25.227
15.6208 | -34.9206 | -25.2349
15.6269 | -34.9211 | -25.243
15.633 | -34.9216 | -25.2511
15.6391 | -34.9221 | -25.2592
15.6452 | -34.9226 | -25.2673
15.6513 | -34.9231 | -25.2751
15.6574 | -34.9236 | -25.2832
15.6635 | -34.9241 | -25.2914
15.6696 | -34.9247 | -25.2995
15.6757 | -34.9252 | -25.3075
15.6818 | -34.9257 | -25.3155
15.688 | -34.9262 | -25.3234
15.6941 | -34.9266 | -25.3315
15.7002 | -34.9271 | -25.3394
15.7063 | -34.9276 | -25.3474
15.7124 | -34.9281 | -25.3555
15.7185 | -34.9286 | -25.3634
15.7246 | -34.9291 | -25.3714
15.7307 | -34.9297 | -25.3794
15.7368 | -34.9302 | -25.3874
15.7429 | -34.9307 | -25.3954
15.749 | -34.9312 | -25.4034
15.7551 | -34.9317 | -25.4115
15.7612 | -34.9322 | -25.4193
15.7673 | -34.9327 | -25.4272
15.7734 | -34.9332 | -25.4352
15.7795 | -34.9338 | -25.4432
15.7856 | -34.9343 | -25.4511
15.7917 | -34.9348 | -25.4587
15.7978 | -34.9353 | -25.4669
15.8039 | -34.9358 | -25.4749
15.81 | -34.9363 | -25.483
15.8161 | -34.9368 | -25.4909
15.8222 | -34.9373 | -25.4988
15.8283 | -34.9378 | -25.5066
15.8345 | -34.9383 | -25.5147
15.8406 | -34.9388 | -25.5227
15.8467 | -34.9394 | -25.5307
15.8528 | -34.9399 | -25.5388
15.8589 | -34.9404 | -25.5468
15.865 | -34.9409 | -25.5545
15.8711 | -34.9414 | -25.5626
15.8772 | -34.9419 | -25.5706
15.8833 | -34.9425 | -25.5787
15.8894 | -34.943 | -25.5867
15.8955 | -34.9434 | -25.5945
15.9016 | -34.944 | -25.6023
15.9077 | -34.9445 | -25.6103
15.9138 | -34.945 | -25.6181
15.9199 | -34.9455 | -25.6261
15.926 | -34.946 | -25.634
15.9321 | -34.9465 | -25.6419
15.9382 | -34.947 | -25.6498
15.9443 | -34.9475 | -25.6577
15.9504 | -34.948 | -25.6658
15.9565 | -34.9485 | -25.6737
15.9626 | -34.949 | -25.6818
15.9687 | -34.9495 | -25.6896
15.9749 | -34.95 | -25.6974
15.981 | -34.9505 | -25.7054
15.9871 | -34.951 | -25.7134
15.9932 | -34.9515 | -25.7213
15.9993 | -34.952 | -25.7293
16.0054 | -34.9525 | -25.7369
16.0115 | -34.953 | -25.7449
16.0176 | -34.9535 | -25.7529
16.0237 | -34.9541 | -25.7609
16.0298 | -34.9545 | -25.7688
16.0359 | -34.955 | -25.7766
16.042 | -34.9556 | -25.7843
16.0481 | -34.9561 | -25.7923
16.0542 | -34.9566 | -25.8002
16.0603 | -34.9571 | -25.808
16.0664 | -34.9576 | -25.8159
16.0725 | -34.9581 | -25.8235
16.0786 | -34.9586 | -25.8313
16.0847 | -34.9591 | -25.8394
16.0908 | -34.9596 | -25.8473
16.0969 | -34.9601 | -25.8554
16.103 | -34.9606 | -25.8633
16.1091 | -34.9612 | -25.8709
16.1152 | -34.9617 | -25.8786
16.1214 | -34.9622 | -25.8865
16.1275 | -34.9627 | -25.8945
16.1336 | -34.9632 | -25.9022
16.1397 | -34.9637 | -25.9101
16.1458 | -34.9642 | -25.9178
16.1519 | -34.9648 | -25.9257
16.158 | -34.9652 | -25.9335
16.1641 | -34.9658 | -25.9415
16.1702 | -34.9663 | -25.9494
16.1763 | -34.9668 | -25.9573
16.1824 | -34.9673 | -25.9649
16.1885 | -34.9678 | -25.9727
16.1946 | -34.9684 | -25.9806
16.2007 | -34.9689 | -25.9885
16.2068 | -34.9694 | -25.9964
16.2129 | -34.9699 | -26.0042
16.219 | -34.9704 | -26.0117
16.2251 | -34.971 | -26.0195
16.2312 | -34.9715 | -26.0272
16.2373 | -34.972 | -26.0351
16.2434 | -34.9725 | -26.0429
16.2495 | -34.973 | -26.0506
16.2556 | -34.9735 | -26.0585
16.2618 | -34.9741 | -26.0663
16.2679 | -34.9746 | -26.0742
16.274 | -34.9751 | -26.082
16.2801 | -34.9756 | -26.0899
16.2862 | -34.9761 | -26.0976
16.2923 | -34.9766 | -26.1055
16.2984 | -34.9771 | -26.1134
16.3045 | -34.9776 | -26.1211
16.3106 | -34.9781 | -26.1289
16.3167 | -34.9787 | -26.1367
16.3228 | -34.9792 | -26.1444
16.3289 | -34.9797 | -26.1522
16.335 | -34.9802 | -26.16
16.3411 | -34.9807 | -26.1678
16.3472 | -34.9812 | -26.1757
16.3533 | -34.9817 | -26.1834
16.3594 | -34.9822 | -26.1909
16.3655 | -34.9827 | -26.1987
16.3716 | -34.9832 | -26.2063
16.3777 | -34.9837 | -26.214
16.3838 | -34.9842 | -26.2218
16.3899 | -34.9848 | -26.2295
16.396 | -34.9853 | -26.2375
16.4021 | -34.9858 | -26.2451
16.4083 | -34.9863 | -26.253
16.4144 | -34.9868 | -26.2607
16.4205 | -34.9873 | -26.2684
16.4266 | -34.9878 | -26.2761
16.4327 | -34.9883 | -26.2838
16.4388 | -34.9888 | -26.2917
16.4449 | -34.9894 | -26.2994
16.451 | -34.9898 | -26.3072
16.4571 | -34.9904 | -26.3151
16.4632 | -34.9909 | -26.3227
16.4693 | -34.9914 | -26.3306
16.4754 | -34.9919 | -26.3384
16.4815 | -34.9924 | -26.3461
16.4876 | -34.9929 | -26.3537
16.4937 | -34.9934 | -26.3614
16.4998 | -34.9939 | -26.3689
16.5059 | -34.9944 | -26.3767
16.512 | -34.9949 | -26.3845
16.5181 | -34.9955 | -26.3923
16.5242 | -34.996 | -26.4001
16.5303 | -34.9965 | -26.4077
16.5364 | -34.9969 | -26.4153
16.5425 | -34.9974 | -26.4232
16.5487 | -34.9979 | -26.4309
16.5548 | -34.9984 | -26.4385
16.5609 | -34.999 | -26.4463
16.567 | -34.9995 | -26.4538
16.5731 | -35 | -26.4617
16.5792 | -35.0005 | -26.4694
16.5853 | -35.001 | -26.4772
16.5914 | -35.0015 | -26.4851
16.5975 | -35.002 | -26.4929
16.6036 | -35.0026 | -26.5005
16.6097 | -35.003 | -26.5083
16.6158 | -35.0035 | -26.516
16.6219 | -35.0041 | -26.5237
16.628 | -35.0046 | -26.5314
16.6341 | -35.0051 | -26.5391
16.6402 | -35.0056 | -26.5467
16.6463 | -35.0061 | -26.5544
16.6524 | -35.0066 | -26.5619
16.6585 | -35.0071 | -26.5697
16.6646 | -35.0077 | -26.5774
16.6707 | -35.0082 | -26.5849
16.6768 | -35.0087 | -26.5926
16.6829 | -35.0092 | -26.6003
16.689 | -35.0097 | -26.6081
16.6952 | -35.0102 | -26.6158
16.7013 | -35.0107 | -26.6236
16.7074 | -35.0112 | -26.631
16.7135 | -35.0118 | -26.6387
16.7196 | -35.0123 | -26.6464
16.7257 | -35.0128 | -26.6541
16.7318 | -35.0133 | -26.6618
16.7379 | -35.0138 | -26.6694
16.744 | -35.0143 | -26.6768
16.7501 | -35.0148 | -26.6846
16.7562 | -35.0153 | -26.6923
16.7623 | -35.0158 | -26.6999
16.7684 | -35.0163 | -26.7077
16.7745 | -35.0169 | -26.7154
16.7806 | -35.0174 | -26.7228
16.7867 | -35.0179 | -26.7304
16.7928 | -35.0184 | -26.7381
16.7989 | -35.0189 | -26.7457
16.805 | -35.0194 | -26.7533
16.8111 | -35.0199 | -26.7608
16.8172 | -35.0204 | -26.7686
16.8233 | -35.0209 | -26.7764
16.8294 | -35.0214 | -26.784
16.8356 | -35.0219 | -26.7918
16.8417 | -35.0225 | -26.7995
16.8478 | -35.023 | -26.8069
16.8539 | -35.0235 | -26.8147
16.86 | -35.024 | -26.8224
16.8661 | -35.0245 | -26.8299
16.8722 | -35.025 | -26.8377
16.8783 | -35.0255 | -26.8454
16.8844 | -35.026 | -26.8529
16.8905 | -35.0265 | -26.8606
16.8966 | -35.027 | -26.8683
16.9027 | -35.0275 | -26.8759
16.9088 | -35.028 | -26.8837
16.9149 | -35.0285 | -26.8914
16.921 | -35.0291 | -26.8987
16.9271 | -35.0295 | -26.9064
16.9332 | -35.03 | -26.914
16.9393 | -35.0305 | -26.9218
16.9454 | -35.031 | -26.9293
16.9515 | -35.0315 | -26.9371
16.9576 | -35.0321 | -26.9444
16.9637 | -35.0326 | -26.9518
16.9698 | -35.0331 | -26.9594
16.9759 | -35.0336 | -26.967
16.9821 | -35.0341 | -26.9747
16.9882 | -35.0346 | -26.9822
16.9943 | -35.0351 | -26.9897
17.0004 | -35.0356 | -26.9974
17.0065 | -35.0362 | -27.0051
17.0126 | -35.0367 | -27.0126
17.0187 | -35.0371 | -27.0203
17.0248 | -35.0377 | -27.0277
17.0309 | -35.0382 | -27.0353
17.037 | -35.0387 | -27.0429
17.0431 | -35.0392 | -27.0504
17.0492 | -35.0397 | -27.0581
17.0553 | -35.0402 | -27.0658
17.0614 | -35.0407 | -27.0733
17.0675 | -35.0412 | -27.081
17.0736 | -35.0417 | -27.0885
17.0797 | -35.0422 | -27.0961
17.0858 | -35.0427 | -27.1038
17.0919 | -35.0432 | -27.1115
17.098 | -35.0437 | -27.1188
17.1041 | -35.0442 | -27.1264
17.1102 | -35.0447 | -27.1339
17.1163 | -35.0452 | -27.1415
17.1225 | -35.0457 | -27.149
17.1286 | -35.0462 | -27.1565
17.1347 | -35.0467 | -27.1642
17.1408 | -35.0473 | -27.1718
17.1469 | -35.0478 | -27.1793
17.153 | -35.0483 | -27.1868
17.1591 | -35.0488 | -27.1945
17.1652 | -35.0493 | -27.2017
17.1713 | -35.0498 | -27.2093
17.1774 | -35.0503 | -27.217
17.1835 | -35.0508 | -27.2245
17.1896 | -35.0513 | -27.232
17.1957 | -35.0518 | -27.2395
17.2018 | -35.0523 | -27.2468
17.2079 | -35.0528 | -27.2544
17.214 | -35.0533 | -27.262
17.2201 | -35.0538 | -27.2695
17.2262 | -35.0543 | -27.2771
17.2323 | -35.0548 | -27.2848
17.2384 | -35.0554 | -27.2924
17.2445 | -35.0558 | -27.2998
17.2506 | -35.0563 | -27.3074
17.2567 | -35.0568 | -27.3151
17.2628 | -35.0573 | -27.3226
17.269 | -35.0578 | -27.3299
17.2751 | -35.0583 | -27.3375
17.2812 | -35.0588 | -27.3451
17.2873 | -35.0594 | -27.3525
17.2934 | -35.0599 | -27.3601
17.2995 | -35.0604 | -27.3678
17.3056 | -35.0609 | -27.3752
17.3117 | -35.0614 | -27.3828
17.3178 | -35.0619 | -27.3903
17.3239 | -35.0624 | -27.3978
17.33 | -35.0629 | -27.4053
17.3361 | -35.0634 | -27.4129
17.3422 | -35.0639 | -27.4203
17.3483 | -35.0644 | -27.4278
17.3544 | -35.0649 | -27.4354
17.3605 | -35.0654 | -27.4429
17.3666 | -35.0659 | -27.4504
17.3727 | -35.0664 | -27.4579
17.3788 | -35.0669 | -27.4653
17.3849 | -35.0674 | -27.4729
17.391 | -35.0679 | -27.4804
17.3971 | -35.0684 | -27.4879
17.4032 | -35.0689 | -27.4954
17.4094 | -35.0694 | -27.5029
17.4155 | -35.0699 | -27.5102
17.4216 | -35.0704 | -27.5177
17.4277 | -35.0709 | -27.5252
17.4338 | -35.0714 | -27.5327
17.4399 | -35.0719 | -27.5402
17.446 | -35.0724 | -27.5475
17.4521 | -35.0729 | -27.5551
17.4582 | -35.0735 | -27.5625
17.4643 | -35.0739 | -27.57
17.4704 | -35.0744 | -27.5775
17.4765 | -35.0749 | -27.5851
17.4826 | -35.0755 | -27.5924
17.4887 | -35.0759 | -27.6
17.4948 | -35.0764 | -27.6075
17.5009 | -35.0769 | -27.615
17.507 | -35.0774 | -27.6225
17.5131 | -35.0779 | -27.6299
17.5192 | -35.0785 | -27.6372
17.5253 | -35.079 | -27.6446
17.5314 | -35.0794 | -27.6521
17.5375 | -35.08 | -27.6596
17.5436 | -35.0805 | -27.6672
17.5497 | -35.081 | -27.6747
17.5559 | -35.0815 | -27.682
17.562 | -35.082 | -27.6894
17.5681 | -35.0825 | -27.6968
17.5742 | -35.083 | -27.7043
17.5803 | -35.0835 | -27.7118
17.5864 | -35.084 | -27.7191
17.5925 | -35.0845 | -27.7267
17.5986 | -35.085 | -27.734
17.6047 | -35.0855 | -27.7414
17.6108 | -35.086 | -27.7489
17.6169 | -35.0865 | -27.7563
17.623 | -35.087 | -27.7636
17.6291 | -35.0875 | -27.771
17.6352 | -35.088 | -27.7783
17.6413 | -35.0885 | -27.7857
17.6474 | -35.089 | -27.7932
17.6535 | -35.0896 | -27.8006
17.6596 | -35.0901 | -27.808
17.6657 | -35.0906 | -27.8156
17.6718 | -35.0911 | -27.823
17.6779 | -35.0916 | -27.8304
17.684 | -35.0921 | -27.8378
17.6901 | -35.0926 | -27.8452
17.6963 | -35.0931 | -27.8525
17.7024 | -35.0936 | -27.8599
17.7085 | -35.0941 | -27.8672
17.7146 | -35.0946 | -27.8746
17.7207 | -35.0951 | -27.8819
17.7268 | -35.0956 | -27.8892
17.7329 | -35.0961 | -27.8966
17.739 | -35.0966 | -27.9041
17.7451 | -35.0971 | -27.9116
17.7512 | -35.0976 | -27.9191
17.7573 | -35.0981 | -27.9266
17.7634 | -35.0986 | -27.9338
17.7695 | -35.0992 | -27.9411
17.7756 | -35.0997 | -27.9485
17.7817 | -35.1002 | -27.956
17.7878 | -35.1007 | -27.9634
17.7939 | -35.1012 | -27.9709
17.8 | -35.1017 | -27.9783
17.8061 | -35.1022 | -27.9856
17.8122 | -35.1027 | -27.993
17.8183 | -35.1032 | -28.0004
17.8244 | -35.1037 | -28.0078
17.8305 | -35.1042 | -28.0151
17.8366 | -35.1047 | -28.0223
17.8428 | -35.1052 | -28.0297
17.8489 | -35.1057 | -28.0371
17.855 | -35.1062 | -28.0444
17.8611 | -35.1067 | -28.0518
17.8672 | -35.1072 | -28.059
17.8733 | -35.1077 | -28.0664
17.8794 | -35.1082 | -28.0738
17.8855 | -35.1087 | -28.0811
17.8916 | -35.1092 | -28.0885
17.8977 | -35.1097 | -28.0959
17.9038 | -35.1102 | -28.1031
17.9099 | -35.1107 | -28.1105
17.916 | -35.1112 | -28.1179
17.9221 | -35.1117 | -28.1253
17.9282 | -35.1122 | -28.1326
17.9343 | -35.1127 | -28.14
17.9404 | -35.1132 | -28.1472
17.9465 | -35.1137 | -28.1545
17.9526 | -35.1142 | -28.162
17.9587 | -35.1147 | -28.1693
17.9648 | -35.1152 | -28.1766
17.9709 | -35.1157 | -28.184
17.977 | -35.1162 | -28.1911
17.9832 | -35.1167 | -28.1985
17.9893 | -35.1172 | -28.206
17.9954 | -35.1177 | -28.2134
18.0015 | -35.1182 | -28.2208
18.0076 | -35.1187 | -28.228
18.0137 | -35.1192 | -28.2354
18.0198 | -35.1197 | -28.2427
18.0259 | -35.1202 | -28.2502
18.032 | -35.1207 | -28.2575
18.0381 | -35.1212 | -28.265
18.0442 | -35.1216 | -28.2723
18.0503 | -35.1221 | -28.2796
18.0564 | -35.1226 | -28.287
18.0625 | -35.1231 | -28.2944
18.0686 | -35.1236 | -28.3018
18.0747 | -35.1241 | -28.3092
18.0808 | -35.1246 | -28.3164
18.0869 | -35.1251 | -28.3236
18.093 | -35.1256 | -28.331
18.0991 | -35.1261 | -28.3384
18.1052 | -35.1266 | -28.3457
18.1113 | -35.1271 | -28.353
18.1174 | -35.1276 | -28.3601
18.1236 | -35.1281 | -28.3676
18.1297 | -35.1286 | -28.375
18.1358 | -35.1291 | -28.3823
18.1419 | -35.1296 | -28.3895
18.148 | -35.1301 | -28.3967
18.1541 | -35.1306 | -28.4039
18.1602 | -35.1311 | -28.4112
18.1663 | -35.1316 | -28.4185
18.1724 | -35.1321 | -28.4259
18.1785 | -35.1326 | -28.4331
18.1846 | -35.1331 | -28.4404
18.1907 | -35.1336 | -28.4477
18.1968 | -35.1341 | -28.4551
18.2029 | -35.1345 | -28.4625
18.209 | -35.135 | -28.4699
18.2151 | -35.1355 | -28.4774
18.2212 | -35.136 | -28.4844
18.2273 | -35.1365 | -28.4917
18.2334 | -35.137 | -28.4991
18.2395 | -35.1375 | -28.5064
18.2456 | -35.138 | -28.5138
18.2517 | -35.1385 | -28.5211
18.2578 | -35.139 | -28.5283
18.2639 | -35.1395 | -28.5357
18.2701 | -35.14 | -28.5429
18.2762 | -35.1405 | -28.5501
18.2823 | -35.141 | -28.5575
18.2884 | -35.1415 | -28.5648
18.2945 | -35.142 | -28.5719
18.3006 | -35.1425 | -28.5793
18.3067 | -35.143 | -28.5867
18.3128 | -35.1435 | -28.594
18.3189 | -35.144 | -28.6014
18.325 | -35.1445 | -28.6084
18.3311 | -35.145 | -28.6158
18.3372 | -35.1455 | -28.6231
18.3433 | -35.1459 | -28.6305
18.3494 | -35.1464 | -28.6378
18.3555 | -35.1469 | -28.645
18.3616 | -35.1475 | -28.652
18.3677 | -35.1479 | -28.6592
18.3738 | -35.1484 | -28.6664
18.3799 | -35.1489 | -28.6736
18.386 | -35.1494 | -28.6808
18.3921 | -35.1499 | -28.688
18.3982 | -35.1503 | -28.6952
18.4043 | -35.1508 | -28.7026
18.4105 | -35.1513 | -28.7099
18.4166 | -35.1518 | -28.7171
18.4227 | -35.1522 | -28.7244
18.4288 | -35.1527 | -28.7316
18.4349 | -35.1532 | -28.7388
18.441 | -35.1536 | -28.746
18.4471 | -35.1541 | -28.7532
18.4532 | -35.1546 | -28.7604
18.4593 | -35.1551 | -28.7676
18.4654 | -35.1556 | -28.7747
18.4715 | -35.1561 | -28.782
18.4776 | -35.1566 | -28.7894
18.4837 | -35.1571 | -28.7967
18.4898 | -35.1575 | -28.804
18.4959 | -35.1581 | -28.8114
18.502 | -35.1586 | -28.8186
18.5081 | -35.1591 | -28.8261
18.5142 | -35.1595 | -28.8334
18.5203 | -35.16 | -28.8405
18.5264 | -35.1605 | -28.8477
18.5325 | -35.161 | -28.855
18.5386 | -35.1615 | -28.862
18.5447 | -35.162 | -28.8693
18.5508 | -35.1625 | -28.8765
18.557 | -35.163 | -28.8838
18.5631 | -35.1635 | -28.8909
18.5692 | -35.1639 | -28.8982
18.5753 | -35.1644 | -28.9051
18.5814 | -35.1649 | -28.9123
18.5875 | -35.1654 | -28.9194
18.5936 | -35.1659 | -28.9268
18.5997 | -35.1664 | -28.934
18.6058 | -35.1668 | -28.9411
18.6119 | -35.1674 | -28.9483
18.618 | -35.1678 | -28.9556
18.6241 | -35.1683 | -28.9628
18.6302 | -35.1688 | -28.9701
18.6363 | -35.1693 | -28.9772
18.6424 | -35.1698 | -28.9843
18.6485 | -35.1703 | -28.9916
18.6546 | -35.1708 | -28.9988
18.6607 | -35.1712 | -29.0061
18.6668 | -35.1717 | -29.0133
18.6729 | -35.1722 | -29.0206
18.679 | -35.1727 | -29.0276
18.6851 | -35.1732 | -29.0347
18.6912 | -35.1737 | -29.042
18.6974 | -35.1742 | -29.0492
18.7035 | -35.1747 | -29.0564
18.7096 | -35.1752 | -29.0636
18.7157 | -35.1757 | -29.0706
18.7218 | -35.1761 | -29.0779
18.7279 | -35.1766 | -29.0851
18.734 | -35.1771 | -29.0923
18.7401 | -35.1777 | -29.0995
18.7462 | -35.1781 | -29.1065
18.7523 | -35.1787 | -29.1135
18.7584 | -35.1791 | -29.1209
18.7645 | -35.1796 | -29.128
18.7706 | -35.1801 | -29.1352
18.7767 | -35.1806 | -29.1425
18.7828 | -35.1811 | -29.1495
18.7889 | -35.1816 | -29.1566
18.795 | -35.182 | -29.1638
18.8011 | -35.1825 | -29.171
18.8072 | -35.183 | -29.1785
18.8133 | -35.1835 | -29.1857
18.8194 | -35.184 | -29.1926
18.8255 | -35.1845 | -29.1998
18.8316 | -35.185 | -29.2071
18.8377 | -35.1855 | -29.2141
18.8439 | -35.186 | -29.2213
18.85 | -35.1864 | -29.2285
18.8561 | -35.1869 | -29.2355
18.8622 | -35.1874 | -29.2428
18.8683 | -35.1879 | -29.25
18.8744 | -35.1884 | -29.2572
18.8805 | -35.1889 | -29.2645
18.8866 | -35.1893 | -29.2718
18.8927 | -35.1898 | -29.2789
18.8988 | -35.1903 | -29.2861
18.9049 | -35.1908 | -29.293
18.911 | -35.1913 | -29.3002
18.9171 | -35.1918 | -29.3075
18.9232 | -35.1923 | -29.3145
18.9293 | -35.1928 | -29.3217
18.9354 | -35.1933 | -29.3289
18.9415 | -35.1937 | -29.336
18.9476 | -35.1942 | -29.3433
18.9537 | -35.1947 | -29.3503
18.9598 | -35.1952 | -29.357
18.9659 | -35.1957 | -29.3642
18.972 | -35.1961 | -29.3714
18.9781 | -35.1966 | -29.3785
18.9843 | -35.1971 | -29.3858
18.9904 | -35.1976 | -29.3929
18.9965 | -35.1981 | -29.3998
19.0026 | -35.1986 | -29.4071
19.0087 | -35.1991 | -29.4142
19.0148 | -35.1996 | -29.4213
19.0209 | -35.2 | -29.4287
19.027 | -35.2005 | -29.4358
19.0331 | -35.201 | -29.4428
19.0392 | -35.2015 | -29.4499
19.0453 | -35.202 | -29.4571
19.0514 | -35.2025 | -29.4642
19.0575 | -35.203 | -29.4715
19.0636 | -35.2034 | -29.4782
19.0697 | -35.2039 | -29.4853
19.0758 | -35.2044 | -29.4925
19.0819 | -35.2049 | -29.4996
19.088 | -35.2053 | -29.5068
19.0941 | -35.2058 | -29.5139
19.1002 | -35.2063 | -29.5209
19.1063 | -35.2067 | -29.5281
19.1124 | -35.2072 | -29.5352
19.1185 | -35.2077 | -29.5423
19.1246 | -35.2082 | -29.5494
19.1308 | -35.2087 | -29.5564
19.1369 | -35.2091 | -29.5633
19.143 | -35.2096 | -29.5705
19.1491 | -35.2101 | -29.5776
19.1552 | -35.2105 | -29.5848
19.1613 | -35.211 | -29.5918
19.1674 | -35.2115 | -29.5988
19.1735 | -35.212 | -29.6058
19.1796 | -35.2125 | -29.6127
19.1857 | -35.2129 | -29.6198
19.1918 | -35.2134 | -29.6269
19.1979 | -35.2139 | -29.6342
19.204 | -35.2144 | -29.6412
19.2101 | -35.2149 | -29.6484
19.2162 | -35.2154 | -29.6556
19.2223 | -35.2158 | -29.6626
19.2284 | -35.2163 | -29.6696
19.2345 | -35.2167 | -29.6768
19.2406 | -35.2173 | -29.6839
19.2467 | -35.2177 | -29.6908
19.2528 | -35.2182 | -29.6982
19.2589 | -35.2187 | -29.7051
19.265 | -35.2192 | -29.7124
19.2712 | -35.2196 | -29.7196
19.2773 | -35.2201 | -29.7266
19.2834 | -35.2206 | -29.7336
19.2895 | -35.2211 | -29.7405
19.2956 | -35.2216 | -29.7477
19.3017 | -35.222 | -29.7546
19.3078 | -35.2225 | -29.7616
19.3139 | -35.2229 | -29.7687
19.32 | -35.2234 | -29.7758
19.3261 | -35.2239 | -29.7828
19.3322 | -35.2244 | -29.7898
19.3383 | -35.2249 | -29.7969
19.3444 | -35.2254 | -29.8038
19.3505 | -35.2258 | -29.8107
19.3566 | -35.2263 | -29.8179
19.3627 | -35.2268 | -29.8249
19.3688 | -35.2272 | -29.8319
19.3749 | -35.2277 | -29.839
19.381 | -35.2282 | -29.8457
19.3871 | -35.2286 | -29.8527
19.3932 | -35.2291 | -29.8596
19.3993 | -35.2296 | -29.8667
19.4054 | -35.23 | -29.8738
19.4115 | -35.2305 | -29.8809
19.4177 | -35.231 | -29.8877
19.4238 | -35.2315 | -29.8947
19.4299 | -35.232 | -29.9017
19.436 | -35.2324 | -29.9089
19.4421 | -35.2329 | -29.9158
19.4482 | -35.2334 | -29.9229
19.4543 | -35.2339 | -29.9297
19.4604 | -35.2343 | -29.9368
19.4665 | -35.2348 | -29.9438
19.4726 | -35.2353 | -29.9508
19.4787 | -35.2357 | -29.9577
19.4848 | -35.2362 | -29.9648
19.4909 | -35.2367 | -29.9717
19.497 | -35.2371 | -29.9785
19.5031 | -35.2376 | -29.9856
19.5092 | -35.238 | -29.9925
19.5153 | -35.2385 | -29.9994
19.5214 | -35.239 | -30.0064
19.5275 | -35.2394 | -30.0135
19.5336 | -35.2399 | -30.0205
19.5397 | -35.2404 | -30.0276
19.5458 | -35.2409 | -30.0347
19.5519 | -35.2414 | -30.0415
19.5581 | -35.2418 | -30.0481
19.5642 | -35.2423 | -30.055
19.5703 | -35.2427 | -30.0622
19.5764 | -35.2432 | -30.0691
19.5825 | -35.2437 | -30.0762
19.5886 | -35.2441 | -30.0833
19.5947 | -35.2446 | -30.09
19.6008 | -35.245 | -30.0969
19.6069 | -35.2455 | -30.1038
19.613 | -35.246 | -30.1109
19.6191 | -35.2465 | -30.1178
19.6252 | -35.2469 | -30.1247
19.6313 | -35.2474 | -30.1313
19.6374 | -35.2479 | -30.1384
19.6435 | -35.2483 | -30.1453
19.6496 | -35.2488 | -30.1522
19.6557 | -35.2492 | -30.1592
19.6618 | -35.2497 | -30.166
19.6679 | -35.2502 | -30.1731
19.674 | -35.2507 | -30.1801
19.6801 | -35.2511 | -30.187
19.6862 | -35.2516 | -30.1941
19.6923 | -35.2521 | -30.2007
19.6984 | -35.2526 | -30.2076
19.7046 | -35.253 | -30.2147
19.7107 | -35.2535 | -30.2215
19.7168 | -35.254 | -30.2282
19.7229 | -35.2544 | -30.2352
19.729 | -35.2549 | -30.2421
19.7351 | -35.2554 | -30.2485
19.7412 | -35.2558 | -30.2553
19.7473 | -35.2563 | -30.2621
19.7534 | -35.2568 | -30.2692
19.7595 | -35.2572 | -30.2759
19.7656 | -35.2577 | -30.2829
19.7717 | -35.2581 | -30.2895
19.7778 | -35.2586 | -30.2963
19.7839 | -35.2591 | -30.3032
19.79 | -35.2595 | -30.3102
19.7961 | -35.26 | -30.3172
19.8022 | -35.2604 | -30.3239
19.8083 | -35.2609 | -30.3311
19.8144 | -35.2614 | -30.3379
19.8205 | -35.2618 | -30.345
19.8266 | -35.2623 | -30.3519
19.8327 | -35.2628 | -30.3588
19.8388 | -35.2632 | -30.3655
19.845 | -35.2637 | -30.3725
19.8511 | -35.2641 | -30.3797
19.8572 | -35.2646 | -30.3865
19.8633 | -35.2651 | -30.3935
19.8694 | -35.2655 | -30.4004
19.8755 | -35.266 | -30.4071
19.8816 | -35.2664 | -30.414
19.8877 | -35.2669 | -30.421
19.8938 | -35.2674 | -30.4278
19.8999 | -35.2678 | -30.4347
19.906 | -35.2682 | -30.4414
19.9121 | -35.2687 | -30.448
19.9182 | -35.2692 | -30.455
19.9243 | -35.2696 | -30.4618
19.9304 | -35.2701 | -30.4686
19.9365 | -35.2705 | -30.4755
19.9426 | -35.271 | -30.4818
19.9487 | -35.2715 | -30.4889
19.9548 | -35.2719 | -30.4959
19.9609 | -35.2724 | -30.5028
19.967 | -35.2729 | -30.5095
19.9731 | -35.2733 | -30.5165
19.9792 | -35.2738 | -30.5233
19.9853 | -35.2742 | -30.5303
19.9915 | -35.2747 | -30.5371
19.9976 | -35.2752 | -30.5438
20.0037 | -35.2756 | -30.5508
20.0098 | -35.276 | -30.5576
20.0159 | -35.2765 | -30.5644
20.022 | -35.277 | -30.5712
20.0281 | -35.2774 | -30.5782
20.0342 | -35.2779 | -30.5851
20.0403 | -35.2783 | -30.5922
20.0464 | -35.2788 | -30.5991
20.0525 | -35.2792 | -30.6059
20.0586 | -35.2797 | -30.6127
20.0647 | -35.2802 | -30.6198
20.0708 | -35.2806 | -30.6268
20.0769 | -35.2811 | -30.6338
20.083 | -35.2815 | -30.6403
20.0891 | -35.282 | -30.6472
20.0952 | -35.2824 | -30.6541
20.1013 | -35.2829 | -30.661
20.1074 | -35.2834 | -30.6681
20.1135 | -35.2838 | -30.6751
20.1196 | -35.2842 | -30.6813
20.1257 | -35.2847 | -30.6883
20.1319 | -35.2851 | -30.6953
20.138 | -35.2856 | -30.702
20.1441 | -35.286 | -30.7089
20.1502 | -35.2865 | -30.7159
20.1563 | -35.2869 | -30.7224
20.1624 | -35.2874 | -30.7293
20.1685 | -35.2878 | -30.7364
20.1746 | -35.2883 | -30.7432
20.1807 | -35.2887 | -30.7501
20.1868 | -35.2892 | -30.7569
20.1929 | -35.2896 | -30.7634
20.199 | -35.2901 | -30.7701
20.2051 | -35.2906 | -30.7767
20.2112 | -35.2911 | -30.7834
20.2173 | -35.2915 | -30.7901
20.2234 | -35.292 | -30.7969
20.2295 | -35.2924 | -30.8037
20.2356 | -35.2929 | -30.8106
20.2417 | -35.2933 | -30.8172
20.2478 | -35.2938 | -30.824
20.2539 | -35.2942 | -30.8308
20.26 | -35.2946 | -30.8375
20.2661 | -35.2951 | -30.8443
20.2723 | -35.2956 | -30.8512
20.2784 | -35.2961 | -30.858
20.2845 | -35.2965 | -30.8647
20.2906 | -35.2969 | -30.8715
20.2967 | -35.2974 | -30.8784
20.3028 | -35.2978 | -30.885
20.3089 | -35.2983 | -30.8917
20.315 | -35.2987 | -30.898
20.3211 | -35.2992 | -30.9048
20.3272 | -35.2997 | -30.9117
20.3333 | -35.3001 | -30.9181
20.3394 | -35.3005 | -30.9248
20.3455 | -35.301 | -30.9317
20.3516 | -35.3014 | -30.9384
20.3577 | -35.3019 | -30.9451
20.3638 | -35.3023 | -30.9519
20.3699 | -35.3028 | -30.9583
20.376 | -35.3032 | -30.9651
20.3821 | -35.3037 | -30.9719
20.3882 | -35.3041 | -30.9788
20.3943 | -35.3046 | -30.9856
20.4004 | -35.305 | -30.9922
20.4065 | -35.3055 | -30.9991
20.4126 | -35.3059 | -31.0056
20.4188 | -35.3064 | -31.0123
20.4249 | -35.3069 | -31.0188
20.431 | -35.3073 | -31.0256
20.4371 | -35.3077 | -31.0318
20.4432 | -35.3082 | -31.0386
20.4493 | -35.3086 | -31.0455
20.4554 | -35.3091 | -31.0522
20.4615 | -35.3096 | -31.0589
20.4676 | -35.31 | -31.0658
20.4737 | -35.3105 | -31.0724
20.4798 | -35.3109 | -31.0789
20.4859 | -35.3113 | -31.0855
20.492 | -35.3118 | -31.0919
20.4981 | -35.3122 | -31.0988
20.5042 | -35.3127 | -31.1055
20.5103 | -35.3131 | -31.1122
20.5164 | -35.3135 | -31.1189
20.5225 | -35.314 | -31.1256
20.5286 | -35.3144 | -31.1325
20.5347 | -35.3149 | -31.1392
20.5408 | -35.3153 | -31.1457
20.5469 | -35.3158 | -31.1524
20.553 | -35.3162 | -31.1594
20.5592 | -35.3167 | -31.1662
20.5653 | -35.3171 | -31.1728
20.5714 | -35.3176 | -31.1792
20.5775 | -35.318 | -31.1855
20.5836 | -35.3185 | -31.1921
20.5897 | -35.3189 | -31.1988
20.5958 | -35.3193 | -31.2055
20.6019 | -35.3198 | -31.2126
20.608 | -35.3202 | -31.2193
20.6141 | -35.3207 | -31.2261
20.6202 | -35.3211 | -31.2328
20.6263 | -35.3216 | -31.2394
20.6324 | -35.322 | -31.2462
20.6385 | -35.3225 | -31.253
20.6446 | -35.3229 | -31.2596
20.6507 | -35.3233 | -31.266
20.6568 | -35.3238 | -31.2729
20.6629 | -35.3243 | -31.28
20.669 | -35.3247 | -31.2867
20.6751 | -35.3252 | -31.2933
20.6812 | -35.3256 | -31.2998
20.6873 | -35.326 | -31.3063
20.6934 | -35.3265 | -31.3129
20.6995 | -35.3269 | -31.3197
20.7057 | -35.3273 | -31.3263
20.7118 | -35.3278 | -31.3329
20.7179 | -35.3282 | -31.3391
20.724 | -35.3287 | -31.346
20.7301 | -35.3291 | -31.3528
20.7362 | -35.3295 | -31.3596
20.7423 | -35.33 | -31.3663
20.7484 | -35.3304 | -31.3732
20.7545 | -35.3309 | -31.3795
20.7606 | -35.3314 | -31.3862
20.7667 | -35.3318 | -31.3929
20.7728 | -35.3323 | -31.3997
20.7789 | -35.3327 | -31.4064
20.785 | -35.3332 | -31.4131
20.7911 | -35.3336 | -31.4198
20.7972 | -35.3341 | -31.4265
20.8033 | -35.3345 | -31.4334
20.8094 | -35.335 | -31.4398
20.8155 | -35.3354 | -31.4466
20.8216 | -35.3359 | -31.4535
20.8277 | -35.3363 | -31.4598
20.8338 | -35.3368 | -31.4666
20.8399 | -35.3373 | -31.4732
20.8461 | -35.3377 | -31.4799
20.8522 | -35.3382 | -31.4868
20.8583 | -35.3387 | -31.4933
20.8644 | -35.3391 | -31.5003
20.8705 | -35.3396 | -31.5071
20.8766 | -35.34 | -31.5137
20.8827 | -35.3405 | -31.5203
20.8888 | -35.3409 | -31.5271
20.8949 | -35.3413 | -31.5334
20.901 | -35.3417 | -31.5401
20.9071 | -35.3421 | -31.5466
20.9132 | -35.3426 | -31.5534
20.9193 | -35.343 | -31.56
20.9254 | -35.3435 | -31.5668
20.9315 | -35.3439 | -31.5734
20.9376 | -35.3443 | -31.5799
20.9437 | -35.3448 | -31.5865
20.9498 | -35.3452 | -31.5929
20.9559 | -35.3456 | -31.5995
20.962 | -35.3461 | -31.6063
20.9681 | -35.3465 | -31.6128
20.9742 | -35.347 | -31.6194
20.9803 | -35.3474 | -31.6261
20.9864 | -35.3479 | -31.6329
20.9926 | -35.3483 | -31.6397
20.9987 | -35.3488 | -31.6461
21.0048 | -35.3492 | -31.6525
21.0109 | -35.3497 | -31.6592
21.017 | -35.3501 | -31.6656
21.0231 | -35.3506 | -31.6723
21.0292 | -35.351 | -31.6788
21.0353 | -35.3515 | -31.6854
21.0414 | -35.3519 | -31.6919
21.0475 | -35.3523 | -31.6986
21.0536 | -35.3528 | -31.7052
21.0597 | -35.3533 | -31.7118
21.0658 | -35.3537 | -31.7183
21.0719 | -35.3541 | -31.7248
21.078 | -35.3546 | -31.7314
21.0841 | -35.355 | -31.7382
21.0902 | -35.3555 | -31.7444
21.0963 | -35.3559 | -31.751
21.1024 | -35.3564 | -31.7577
21.1085 | -35.3568 | -31.7644
21.1146 | -35.3572 | -31.771
21.1207 | -35.3576 | -31.7776
21.1268 | -35.3581 | -31.7845
21.133 | -35.3585 | -31.791
21.1391 | -35.3589 | -31.7974
21.1452 | -35.3594 | -31.8038
21.1513 | -35.3598 | -31.8105
21.1574 | -35.3601 | -31.8171
21.1635 | -35.3606 | -31.8234
21.1696 | -35.3611 | -31.8299
21.1757 | -35.3615 | -31.8364
21.1818 | -35.3619 | -31.8428
21.1879 | -35.3624 | -31.8491
21.194 | -35.3628 | -31.8552
21.2001 | -35.3632 | -31.8621
21.2062 | -35.3636 | -31.8687
21.2123 | -35.3641 | -31.8747
21.2184 | -35.3646 | -31.8812
21.2245 | -35.3651 | -31.8879
21.2306 | -35.3655 | -31.8942
21.2367 | -35.366 | -31.9008
21.2428 | -35.3664 | -31.9071
21.2489 | -35.3668 | -31.9136
21.255 | -35.3672 | -31.9202
21.2611 | -35.3677 | -31.9268
21.2672 | -35.3681 | -31.9334
21.2733 | -35.3685 | -31.9402
21.2795 | -35.3689 | -31.9466
21.2856 | -35.3694 | -31.9532
21.2917 | -35.3699 | -31.9597
21.2978 | -35.3703 | -31.9664
21.3039 | -35.3708 | -31.9732
21.31 | -35.3713 | -31.9799
21.3161 | -35.3717 | -31.9865
21.3222 | -35.3721 | -31.9931
21.3283 | -35.3726 | -31.9994
21.3344 | -35.373 | -32.0062
21.3405 | -35.3735 | -32.0131
21.3466 | -35.3739 | -32.0196
21.3527 | -35.3743 | -32.0259
21.3588 | -35.3747 | -32.0323
21.3649 | -35.3751 | -32.0387
21.371 | -35.3756 | -32.0451
21.3771 | -35.376 | -32.0514
21.3832 | -35.3765 | -32.0579
21.3893 | -35.3769 | -32.0638
21.3954 | -35.3773 | -32.0701
21.4015 | -35.3776 | -32.0764
21.4076 | -35.3781 | -32.0832
21.4137 | -35.3786 | -32.0897
21.4199 | -35.379 | -32.0961
21.426 | -35.3794 | -32.1027
21.4321 | -35.3799 | -32.1091
21.4382 | -35.3803 | -32.1156
21.4443 | -35.3808 | -32.1218
21.4504 | -35.3812 | -32.1286
21.4565 | -35.3816 | -32.1344
21.4626 | -35.382 | -32.1409
21.4687 | -35.3824 | -32.1472
21.4748 | -35.3829 | -32.1534
21.4809 | -35.3833 | -32.16
21.487 | -35.3837 | -32.1664
21.4931 | -35.3842 | -32.1726
21.4992 | -35.3846 | -32.1789
21.5053 | -35.385 | -32.185
21.5114 | -35.3854 | -32.1912
21.5175 | -35.3859 | -32.1973
21.5236 | -35.3863 | -32.2041
21.5297 | -35.3868 | -32.2104
21.5358 | -35.3872 | -32.2167
21.5419 | -35.3877 | -32.2238
21.548 | -35.3881 | -32.2299
21.5541 | -35.3885 | -32.2368
21.5602 | -35.389 | -32.243
21.5664 | -35.3895 | -32.2494
21.5725 | -35.3899 | -32.2559
21.5786 | -35.3903 | -32.2624
21.5847 | -35.3908 | -32.269
21.5908 | -35.3912 | -32.2753
21.5969 | -35.3916 | -32.282
21.603 | -35.392 | -32.2884
21.6091 | -35.3925 | -32.2949
21.6152 | -35.3929 | -32.301
21.6213 | -35.3933 | -32.3076
21.6274 | -35.3938 | -32.3141
21.6335 | -35.3942 | -32.3206
21.6396 | -35.3946 | -32.3271
21.6457 | -35.3951 | -32.3333
21.6518 | -35.3955 | -32.3395
21.6579 | -35.3959 | -32.3461
21.664 | -35.3963 | -32.3529
21.6701 | -35.3967 | -32.3592
21.6762 | -35.3971 | -32.3654
21.6823 | -35.3975 | -32.3723
21.6884 | -35.3979 | -32.3785
21.6945 | -35.3983 | -32.385
21.7006 | -35.3987 | -32.3914
21.7068 | -35.3992 | -32.3975
21.7129 | -35.3996 | -32.4042
21.719 | -35.4001 | -32.4106
21.7251 | -35.4006 | -32.4171
21.7312 | -35.401 | -32.4236
21.7373 | -35.4013 | -32.4298
21.7434 | -35.4018 | -32.436
21.7495 | -35.4022 | -32.4426
21.7556 | -35.4026 | -32.4491
21.7617 | -35.403 | -32.4555
21.7678 | -35.4034 | -32.4616
21.7739 | -35.4039 | -32.468
21.78 | -35.4043 | -32.4744
21.7861 | -35.4047 | -32.4809
21.7922 | -35.4051 | -32.4873
21.7983 | -35.4055 | -32.4935
21.8044 | -35.406 | -32.5
21.8105 | -35.4064 | -32.5068
21.8166 | -35.4069 | -32.513
21.8227 | -35.4073 | -32.5194
21.8288 | -35.4077 | -32.5256
21.8349 | -35.4081 | -32.5322
21.841 | -35.4085 | -32.5388
21.8471 | -35.409 | -32.545
21.8533 | -35.4095 | -32.551
21.8594 | -35.4099 | -32.5575
21.8655 | -35.4104 | -32.5636
21.8716 | -35.4108 | -32.5703
21.8777 | -35.4112 | -32.5764
21.8838 | -35.4117 | -32.5825
21.8899 | -35.4122 | -32.5892
21.896 | -35.4126 | -32.5957
21.9021 | -35.4131 | -32.6022
21.9082 | -35.4135 | -32.6084
21.9143 | -35.414 | -32.615
21.9204 | -35.4144 | -32.6207
21.9265 | -35.4148 | -32.6272
21.9326 | -35.4152 | -32.6337
21.9387 | -35.4157 | -32.64
21.9448 | -35.4161 | -32.6462
21.9509 | -35.4166 | -32.6524
21.957 | -35.417 | -32.659
21.9631 | -35.4174 | -32.665
21.9692 | -35.4178 | -32.6715
21.9753 | -35.4182 | -32.678
21.9814 | -35.4187 | -32.6844
21.9875 | -35.4192 | -32.6908
21.9937 | -35.4196 | -32.6968
21.9998 | -35.42 | -32.7032
22.0059 | -35.4205 | -32.7093
22.012 | -35.4209 | -32.7158
22.0181 | -35.4215 | -32.7214
22.0242 | -35.4219 | -32.7281
22.0303 | -35.4223 | -32.735
22.0364 | -35.4227 | -32.7414
22.0425 | -35.4231 | -32.7475
22.0486 | -35.4236 | -32.7541
22.0547 | -35.4241 | -32.7605
22.0608 | -35.4245 | -32.7663
22.0669 | -35.4249 | -32.7724
22.073 | -35.4253 | -32.7787
22.0791 | -35.4258 | -32.7853
22.0852 | -35.4263 | -32.792
22.0913 | -35.4267 | -32.7982
22.0974 | -35.4271 | -32.8046
22.1035 | -35.4276 | -32.8108
22.1096 | -35.428 | -32.817
22.1157 | -35.4284 | -32.8239
22.1218 | -35.4288 | -32.8299
22.1279 | -35.4293 | -32.8359
22.134 | -35.4297 | -32.8422
22.1402 | -35.4302 | -32.8488
22.1463 | -35.4306 | -32.8551
22.1524 | -35.431 | -32.8618
22.1585 | -35.4314 | -32.8685
22.1646 | -35.4319 | -32.8742
22.1707 | -35.4323 | -32.8806
22.1768 | -35.4328 | -32.8868
22.1829 | -35.4332 | -32.8929
22.189 | -35.4337 | -32.899
22.1951 | -35.4341 | -32.905
22.2012 | -35.4345 | -32.9115
22.2073 | -35.435 | -32.9177
22.2134 | -35.4354 | -32.9238
22.2195 | -35.4358 | -32.9299
22.2256 | -35.4363 | -32.9362
22.2317 | -35.4368 | -32.9421
22.2378 | -35.4373 | -32.9482
22.2439 | -35.4377 | -32.9545
22.25 | -35.4381 | -32.9608
22.2561 | -35.4385 | -32.9671
22.2622 | -35.4389 | -32.9736
22.2683 | -35.4394 | -32.9802
22.2744 | -35.4398 | -32.9868
22.2806 | -35.4403 | -32.9931
22.2867 | -35.4407 | -32.9994
22.2928 | -35.4412 | -33.0057
22.2989 | -35.4417 | -33.0121
22.305 | -35.4421 | -33.0181
22.3111 | -35.4426 | -33.0244
22.3172 | -35.4431 | -33.0305
22.3233 | -35.4435 | -33.0369
22.3294 | -35.444 | -33.0434
22.3355 | -35.4444 | -33.0491
22.3416 | -35.4449 | -33.0555
22.3477 | -35.4453 | -33.0617
22.3538 | -35.4458 | -33.0682
22.3599 | -35.4462 | -33.0744
22.366 | -35.4466 | -33.0806
22.3721 | -35.447 | -33.0871
22.3782 | -35.4473 | -33.0932
22.3843 | -35.4477 | -33.0998
22.3904 | -35.4481 | -33.106
22.3965 | -35.4485 | -33.1125
22.4026 | -35.449 | -33.1186
22.4087 | -35.4494 | -33.1248
22.4148 | -35.4499 | -33.1313
22.4209 | -35.4503 | -33.1376
22.4271 | -35.4507 | -33.1439
22.4332 | -35.4511 | -33.1503
22.4393 | -35.4515 | -33.1564
22.4454 | -35.452 | -33.1627
22.4515 | -35.4524 | -33.1689
22.4576 | -35.4528 | -33.1751
22.4637 | -35.4533 | -33.1812
22.4698 | -35.4537 | -33.1874
22.4759 | -35.4542 | -33.1936
22.482 | -35.4546 | -33.1999
22.4881 | -35.455 | -33.2065
22.4942 | -35.4554 | -33.2124
22.5003 | -35.4558 | -33.2192
22.5064 | -35.4562 | -33.2259
22.5125 | -35.4566 | -33.2319
22.5186 | -35.4571 | -33.238
22.5247 | -35.4575 | -33.2441
22.5308 | -35.4579 | -33.2506
22.5369 | -35.4584 | -33.2569
22.543 | -35.4588 | -33.2632
22.5491 | -35.4592 | -33.2693
22.5552 | -35.4596 | -33.2756
22.5613 | -35.4601 | -33.2819
22.5675 | -35.4605 | -33.288
22.5736 | -35.4609 | -33.2943
22.5797 | -35.4613 | -33.3009
22.5858 | -35.4617 | -33.3067
22.5919 | -35.4621 | -33.3129
22.598 | -35.4625 | -33.3193
22.6041 | -35.463 | -33.3256
22.6102 | -35.4634 | -33.3321
22.6163 | -35.4638 | -33.3379
22.6224 | -35.4642 | -33.3439
22.6285 | -35.4647 | -33.3504
22.6346 | -35.4651 | -33.3567
22.6407 | -35.4655 | -33.3627
22.6468 | -35.466 | -33.3692
22.6529 | -35.4664 | -33.3754
22.659 | -35.4668 | -33.3816
22.6651 | -35.4673 | -33.3879
22.6712 | -35.4677 | -33.3943
22.6773 | -35.4681 | -33.4004
22.6834 | -35.4686 | -33.4067
22.6895 | -35.469 | -33.4125
22.6956 | -35.4694 | -33.4182
22.7017 | -35.4699 | -33.4242
22.7079 | -35.4703 | -33.4305
22.714 | -35.4708 | -33.4366
22.7201 | -35.4712 | -33.4429
22.7262 | -35.4717 | -33.4489
22.7323 | -35.4721 | -33.4549
22.7384 | -35.4726 | -33.4612
22.7445 | -35.473 | -33.4677
22.7506 | -35.4735 | -33.4741
22.7567 | -35.4739 | -33.4802
22.7628 | -35.4744 | -33.4859
22.7689 | -35.4748 | -33.4923
22.775 | -35.4753 | -33.4987
22.7811 | -35.4757 | -33.505
22.7872 | -35.4762 | -33.5113
22.7933 | -35.4765 | -33.5172
22.7994 | -35.477 | -33.5236
22.8055 | -35.4774 | -33.5299
22.8116 | -35.4778 | -33.5361
22.8177 | -35.4783 | -33.5422
22.8238 | -35.4787 | -33.5484
22.8299 | -35.4792 | -33.5546
22.836 | -35.4796 | -33.561
22.8421 | -35.48 | -33.5672
22.8482 | -35.4804 | -33.5733
22.8544 | -35.4809 | -33.5792
22.8605 | -35.4813 | -33.5855
22.8666 | -35.4817 | -33.5914
22.8727 | -35.4822 | -33.5975
22.8788 | -35.4826 | -33.6035
22.8849 | -35.483 | -33.6099
22.891 | -35.4835 | -33.6162
22.8971 | -35.4839 | -33.6226
22.9032 | -35.4843 | -33.6284
22.9093 | -35.4847 | -33.635
22.9154 | -35.4851 | -33.6412
22.9215 | -35.4856 | -33.6471
22.9276 | -35.486 | -33.6536
22.9337 | -35.4865 | -33.6594
22.9398 | -35.4868 | -33.6655
22.9459 | -35.4873 | -33.6719
22.952 | -35.4877 | -33.6782
22.9581 | -35.4882 | -33.6846
22.9642 | -35.4886 | -33.6906
22.9703 | -35.4891 | -33.6968
22.9764 | -35.4895 | -33.703
22.9825 | -35.4899 | -33.709
22.9886 | -35.4903 | -33.7151
22.9948 | -35.4907 | -33.7211
23.0009 | -35.4911 | -33.7275
23.007 | -35.4915 | -33.7335
23.0131 | -35.4919 | -33.7397
23.0192 | -35.4924 | -33.7461
23.0253 | -35.4928 | -33.7523
23.0314 | -35.4932 | -33.7585
23.0375 | -35.4936 | -33.7646
23.0436 | -35.494 | -33.7705
23.0497 | -35.4945 | -33.7765
23.0558 | -35.495 | -33.7826
23.0619 | -35.4954 | -33.7885
23.068 | -35.4959 | -33.7948
23.0741 | -35.4963 | -33.8005
23.0802 | -35.4968 | -33.8067
23.0863 | -35.4972 | -33.8128
23.0924 | -35.4976 | -33.819
23.0985 | -35.4981 | -33.825
23.1046 | -35.4985 | -33.8314
23.1107 | -35.499 | -33.8374
23.1168 | -35.4994 | -33.8435
23.1229 | -35.4998 | -33.8492
23.129 | -35.5002 | -33.8554
23.1351 | -35.5007 | -33.8611
23.1413 | -35.5012 | -33.8676
23.1474 | -35.5015 | -33.8733
23.1535 | -35.5019 | -33.8797
23.1596 | -35.5024 | -33.8856
23.1657 | -35.5028 | -33.892
23.1718 | -35.5033 | -33.8981
23.1779 | -35.5037 | -33.9046
23.184 | -35.5041 | -33.9103
23.1901 | -35.5045 | -33.9165
23.1962 | -35.505 | -33.9226
23.2023 | -35.5054 | -33.9287
23.2084 | -35.5059 | -33.9348
23.2145 | -35.5063 | -33.9404
23.2206 | -35.5068 | -33.9465
23.2267 | -35.5072 | -33.9528
23.2328 | -35.5076 | -33.9591
23.2389 | -35.5081 | -33.9651
23.245 | -35.5085 | -33.9713
23.2511 | -35.5089 | -33.9774
23.2572 | -35.5094 | -33.9835
23.2633 | -35.5098 | -33.9895
23.2694 | -35.5103 | -33.9955
23.2755 | -35.5107 | -34.0019
23.2817 | -35.5111 | -34.0078
23.2878 | -35.5114 | -34.0135
23.2939 | -35.5119 | -34.0193
23.3 | -35.5123 | -34.0254
23.3061 | -35.5127 | -34.0314
23.3122 | -35.5131 | -34.0376
23.3183 | -35.5136 | -34.0434
23.3244 | -35.514 | -34.0494
23.3305 | -35.5144 | -34.0558
23.3366 | -35.5149 | -34.0617
23.3427 | -35.5153 | -34.0677
23.3488 | -35.5157 | -34.074
23.3549 | -35.5161 | -34.0802
23.361 | -35.5166 | -34.0861
23.3671 | -35.517 | -34.0922
23.3732 | -35.5175 | -34.0982
23.3793 | -35.5179 | -34.1042
23.3854 | -35.5183 | -34.1103
23.3915 | -35.5187 | -34.1161
23.3976 | -35.5191 | -34.1219
23.4037 | -35.5196 | -34.1278
23.4098 | -35.52 | -34.1337
23.4159 | -35.5204 | -34.1399
23.422 | -35.5209 | -34.1462
23.4282 | -35.5213 | -34.1518
23.4343 | -35.5217 | -34.1577
23.4404 | -35.5222 | -34.1639
23.4465 | -35.5226 | -34.1703
23.4526 | -35.523 | -34.1764
23.4587 | -35.5234 | -34.1825
23.4648 | -35.5238 | -34.1881
23.4709 | -35.5242 | -34.1942
23.477 | -35.5246 | -34.2004
23.4831 | -35.525 | -34.2066
23.4892 | -35.5255 | -34.2126
23.4953 | -35.5259 | -34.2188
23.5014 | -35.5263 | -34.2249
23.5075 | -35.5268 | -34.231
23.5136 | -35.5272 | -34.2371
23.5197 | -35.5277 | -34.2433
23.5258 | -35.5281 | -34.2491
23.5319 | -35.5285 | -34.2548
23.538 | -35.5289 | -34.261
23.5441 | -35.5294 | -34.2671
23.5502 | -35.5298 | -34.273
23.5563 | -35.5303 | -34.2791
23.5624 | -35.5307 | -34.285
23.5686 | -35.5312 | -34.2911
23.5747 | -35.5316 | -34.2972
23.5808 | -35.532 | -34.3032
23.5869 | -35.5324 | -34.3091
23.593 | -35.5329 | -34.3152
23.5991 | -35.5333 | -34.321
23.6052 | -35.5337 | -34.327
23.6113 | -35.5341 | -34.3334
23.6174 | -35.5345 | -34.3395
23.6235 | -35.5349 | -34.3459
23.6296 | -35.5353 | -34.3521
23.6357 | -35.5358 | -34.3583
23.6418 | -35.5362 | -34.364
23.6479 | -35.5366 | -34.3697
23.654 | -35.537 | -34.376
23.6601 | -35.5375 | -34.3819
23.6662 | -35.5379 | -34.388
23.6723 | -35.5384 | -34.3938
23.6784 | -35.5388 | -34.3997
23.6845 | -35.5392 | -34.406
23.6906 | -35.5396 | -34.4121
23.6967 | -35.54 | -34.4179
23.7028 | -35.5405 | -34.4239
23.7089 | -35.5409 | -34.4295
23.7151 | -35.5413 | -34.4354
23.7212 | -35.5418 | -34.4415
23.7273 | -35.5422 | -34.4475
23.7334 | -35.5426 | -34.4532
23.7395 | -35.543 | -34.4594
23.7456 | -35.5434 | -34.4655
23.7517 | -35.5438 | -34.4712
23.7578 | -35.5442 | -34.4771
23.7639 | -35.5446 | -34.4831
23.77 | -35.5451 | -34.4891
23.7761 | -35.5455 | -34.4951
23.7822 | -35.546 | -34.501
23.7883 | -35.5464 | -34.507
23.7944 | -35.5468 | -34.5128
23.8005 | -35.5472 | -34.5189
23.8066 | -35.5477 | -34.5249
23.8127 | -35.5481 | -34.5306
23.8188 | -35.5486 | -34.5369
23.8249 | -35.549 | -34.5428
23.831 | -35.5494 | -34.5487
23.8371 | -35.5498 | -34.5546
23.8432 | -35.5502 | -34.5608
23.8493 | -35.5507 | -34.5668
23.8555 | -35.5511 | -34.573
23.8616 | -35.5516 | -34.5789
23.8677 | -35.552 | -34.5849
23.8738 | -35.5524 | -34.5909
23.8799 | -35.5528 | -34.5968
23.886 | -35.5532 | -34.6025
23.8921 | -35.5536 | -34.6084
23.8982 | -35.5541 | -34.6142
23.9043 | -35.5545 | -34.6202
23.9104 | -35.5549 | -34.6262
23.9165 | -35.5553 | -34.6321
23.9226 | -35.5558 | -34.6381
23.9287 | -35.5562 | -34.644
23.9348 | -35.5566 | -34.6501
23.9409 | -35.557 | -34.6562
23.947 | -35.5574 | -34.662
23.9531 | -35.5579 | -34.6679
23.9592 | -35.5583 | -34.6742
23.9653 | -35.5587 | -34.6801
23.9714 | -35.5592 | -34.686
23.9775 | -35.5596 | -34.6919
23.9836 | -35.56 | -34.6981
23.9897 | -35.5604 | -34.7043
23.9958 | -35.5609 | -34.7103
24.002 | -35.5613 | -34.7161
24.0081 | -35.5617 | -34.7221
24.0142 | -35.5621 | -34.7282
24.0203 | -35.5624 | -34.7341
24.0264 | -35.5628 | -34.7395
24.0325 | -35.5632 | -34.7456
24.0386 | -35.5637 | -34.7521
24.0447 | -35.5641 | -34.7582
24.0508 | -35.5645 | -34.764
24.0569 | -35.565 | -34.7699
24.063 | -35.5654 | -34.7758
24.0691 | -35.5658 | -34.7816
24.0752 | -35.5663 | -34.7877
24.0813 | -35.5667 | -34.7934
24.0874 | -35.5671 | -34.7991
24.0935 | -35.5675 | -34.8051
24.0996 | -35.5679 | -34.811
24.1057 | -35.5684 | -34.8172
24.1118 | -35.5688 | -34.8233
24.1179 | -35.5692 | -34.8294
24.124 | -35.5696 | -34.8353
24.1301 | -35.5701 | -34.8417
24.1362 | -35.5705 | -34.8479
24.1424 | -35.5709 | -34.8539
24.1485 | -35.5713 | -34.86
24.1546 | -35.5717 | -34.866
24.1607 | -35.5721 | -34.8721
24.1668 | -35.5725 | -34.8778
24.1729 | -35.5729 | -34.8841
24.179 | -35.5733 | -34.8902
24.1851 | -35.5737 | -34.8962
24.1912 | -35.5742 | -34.9021
24.1973 | -35.5746 | -34.908
24.2034 | -35.5751 | -34.9139
24.2095 | -35.5756 | -34.9198
24.2156 | -35.576 | -34.9257
24.2217 | -35.5764 | -34.9314
24.2278 | -35.5768 | -34.9372
24.2339 | -35.5772 | -34.9434
24.24 | -35.5776 | -34.9488
24.2461 | -35.578 | -34.9546
24.2522 | -35.5784 | -34.9604
24.2583 | -35.5788 | -34.9666
24.2644 | -35.5793 | -34.9727
24.2705 | -35.5797 | -34.9785
24.2766 | -35.5802 | -34.9844
24.2827 | -35.5806 | -34.9903
24.2889 | -35.581 | -34.9965
24.295 | -35.5815 | -35.0022
24.3011 | -35.5819 | -35.008
24.3072 | -35.5822 | -35.0134
24.3133 | -35.5826 | -35.0197
24.3194 | -35.583 | -35.0258
24.3255 | -35.5834 | -35.0319
24.3316 | -35.5838 | -35.0378
24.3377 | -35.5842 | -35.0436
24.3438 | -35.5846 | -35.0497
24.3499 | -35.5851 | -35.0554
24.356 | -35.5855 | -35.0613
24.3621 | -35.5859 | -35.067
24.3682 | -35.5863 | -35.0729
24.3743 | -35.5867 | -35.079
24.3804 | -35.587 | -35.0848
24.3865 | -35.5875 | -35.091
24.3926 | -35.5879 | -35.0971
24.3987 | -35.5883 | -35.1029
24.4048 | -35.5887 | -35.1085
24.4109 | -35.5891 | -35.1143
24.417 | -35.5896 | -35.12
24.4231 | -35.59 | -35.1262
24.4293 | -35.5904 | -35.1321
24.4354 | -35.5908 | -35.1381
24.4415 | -35.5912 | -35.144
24.4476 | -35.5917 | -35.1496
24.4537 | -35.5921 | -35.1552
24.4598 | -35.5925 | -35.1613
24.4659 | -35.5929 | -35.1675
24.472 | -35.5934 | -35.1733
24.4781 | -35.5938 | -35.1793
24.4842 | -35.5942 | -35.1853
24.4903 | -35.5946 | -35.1911
24.4964 | -35.595 | -35.197
24.5025 | -35.5955 | -35.2027
24.5086 | -35.5959 | -35.2085
24.5147 | -35.5963 | -35.2142
24.5208 | -35.5967 | -35.2197
24.5269 | -35.5971 | -35.2255
24.533 | -35.5975 | -35.2316
24.5391 | -35.5979 | -35.2375
24.5452 | -35.5983 | -35.2431
24.5513 | -35.5987 | -35.2488
24.5574 | -35.5991 | -35.2547
24.5635 | -35.5996 | -35.2606
24.5696 | -35.6 | -35.2665
24.5758 | -35.6004 | -35.2722
24.5819 | -35.6008 | -35.2778
24.588 | -35.6012 | -35.2834
24.5941 | -35.6016 | -35.2893
24.6002 | -35.602 | -35.2951
24.6063 | -35.6024 | -35.3006
24.6124 | -35.6028 | -35.3067
24.6185 | -35.6033 | -35.3124
24.6246 | -35.6037 | -35.3184
24.6307 | -35.604 | -35.3244
24.6368 | -35.6045 | -35.3303
24.6429 | -35.6049 | -35.3361
24.649 | -35.6053 | -35.3419
24.6551 | -35.6057 | -35.3479
24.6612 | -35.6061 | -35.3536
24.6673 | -35.6065 | -35.3593
24.6734 | -35.6069 | -35.3652
24.6795 | -35.6073 | -35.3711
24.6856 | -35.6077 | -35.3772
24.6917 | -35.6082 | -35.383
24.6978 | -35.6086 | -35.3887
24.7039 | -35.609 | -35.3946
24.71 | -35.6094 | -35.4005
24.7162 | -35.6098 | -35.4065
24.7223 | -35.6102 | -35.4123
24.7284 | -35.6106 | -35.4177
24.7345 | -35.611 | -35.4236
24.7406 | -35.6114 | -35.4295
24.7467 | -35.6118 | -35.4354
24.7528 | -35.6123 | -35.4413
24.7589 | -35.6127 | -35.447
24.765 | -35.6131 | -35.4527
24.7711 | -35.6135 | -35.4586
24.7772 | -35.6139 | -35.4645
24.7833 | -35.6143 | -35.4702
24.7894 | -35.6147 | -35.4763
24.7955 | -35.6151 | -35.4825
24.8016 | -35.6155 | -35.4884
24.8077 | -35.6159 | -35.4943
24.8138 | -35.6163 | -35.5003
24.8199 | -35.6167 | -35.5062
24.826 | -35.6171 | -35.5119
24.8321 | -35.6175 | -35.5178
24.8382 | -35.6179 | -35.5236
24.8443 | -35.6183 | -35.5295
24.8504 | -35.6187 | -35.5353
24.8565 | -35.6192 | -35.5412
24.8627 | -35.6196 | -35.5468
24.8688 | -35.62 | -35.5525
24.8749 | -35.6204 | -35.5582
24.881 | -35.6209 | -35.5641
24.8871 | -35.6213 | -35.5698
24.8932 | -35.6217 | -35.5758
24.8993 | -35.6222 | -35.5817
24.9054 | -35.6226 | -35.5872
24.9115 | -35.623 | -35.5929
24.9176 | -35.6234 | -35.5988
24.9237 | -35.6238 | -35.6045
24.9298 | -35.6241 | -35.6101
24.9359 | -35.6246 | -35.6161
24.942 | -35.625 | -35.622
24.9481 | -35.6254 | -35.628
24.9542 | -35.6258 | -35.6337
24.9603 | -35.6262 | -35.6395
24.9664 | -35.6266 | -35.6451
24.9725 | -35.627 | -35.6508
24.9786 | -35.6274 | -35.6561
24.9847 | -35.6278 | -35.6618
24.9908 | -35.6283 | -35.6677
24.9969 | -35.6287 | -35.6736
25.0031 | -35.6291 | -35.6794
25.0092 | -35.6295 | -35.6849
25.0153 | -35.6299 | -35.6906
25.0214 | -35.6303 | -35.6965
25.0275 | -35.6307 | -35.7023
25.0336 | -35.6312 | -35.7081
25.0397 | -35.6316 | -35.7138
25.0458 | -35.632 | -35.7195
25.0519 | -35.6324 | -35.7252
25.058 | -35.6328 | -35.7311
25.0641 | -35.6332 | -35.7368
25.0702 | -35.6336 | -35.7425
25.0763 | -35.6341 | -35.7482
25.0824 | -35.6344 | -35.7537
25.0885 | -35.6348 | -35.7595
25.0946 | -35.6352 | -35.765
25.1007 | -35.6356 | -35.7706
25.1068 | -35.636 | -35.7766
25.1129 | -35.6364 | -35.7825
25.119 | -35.6369 | -35.7881
25.1251 | -35.6373 | -35.7939
25.1312 | -35.6377 | -35.7997
25.1373 | -35.6381 | -35.8055
25.1435 | -35.6385 | -35.8114
25.1496 | -35.639 | -35.8169
25.1557 | -35.6394 | -35.8228
25.1618 | -35.6398 | -35.8286
25.1679 | -35.6402 | -35.8343
25.174 | -35.6406 | -35.84
25.1801 | -35.641 | -35.846
25.1862 | -35.6414 | -35.8515
25.1923 | -35.6418 | -35.8571
25.1984 | -35.6423 | -35.8628
25.2045 | -35.6427 | -35.8683
25.2106 | -35.6431 | -35.874
25.2167 | -35.6435 | -35.8798
25.2228 | -35.6438 | -35.8856
25.2289 | -35.6443 | -35.8912
25.235 | -35.6447 | -35.8969
25.2411 | -35.6451 | -35.9026
25.2472 | -35.6455 | -35.9086
25.2533 | -35.6459 | -35.9141
25.2594 | -35.6463 | -35.9195
25.2655 | -35.6466 | -35.9255
25.2716 | -35.647 | -35.9312
25.2777 | -35.6474 | -35.9371
25.2838 | -35.6478 | -35.9427
25.29 | -35.6483 | -35.9483
25.2961 | -35.6487 | -35.9536
25.3022 | -35.6491 | -35.9594
25.3083 | -35.6495 | -35.9654
25.3144 | -35.6499 | -35.9711
25.3205 | -35.6502 | -35.9768
25.3266 | -35.6506 | -35.9824
25.3327 | -35.651 | -35.9884
25.3388 | -35.6514 | -35.994
25.3449 | -35.6518 | -35.9998
25.351 | -35.6522 | -36.0056
25.3571 | -35.6525 | -36.0114
25.3632 | -35.6529 | -36.0164
25.3693 | -35.6533 | -36.0225
25.3754 | -35.6537 | -36.0281
25.3815 | -35.6541 | -36.0337
25.3876 | -35.6546 | -36.0395
25.3937 | -35.655 | -36.0453
25.3998 | -35.6553 | -36.0507
25.4059 | -35.6558 | -36.0564
25.412 | -35.6562 | -36.0619
25.4181 | -35.6566 | -36.0679
25.4242 | -35.657 | -36.0738
25.4304 | -35.6574 | -36.0798
25.4365 | -35.6578 | -36.0852
25.4426 | -35.6582 | -36.091
25.4487 | -35.6585 | -36.0969
25.4548 | -35.6589 | -36.1029
25.4609 | -35.6593 | -36.1084
25.467 | -35.6598 | -36.1141
25.4731 | -35.6601 | -36.1195
25.4792 | -35.6605 | -36.1252
25.4853 | -35.6609 | -36.131
25.4914 | -35.6613 | -36.1366
25.4975 | -35.6617 | -36.1421
25.5036 | -35.6622 | -36.1477
25.5097 | -35.6626 | -36.1531
25.5158 | -35.663 | -36.1587
25.5219 | -35.6634 | -36.1644
25.528 | -35.6638 | -36.17
25.5341 | -35.6642 | -36.176
25.5402 | -35.6645 | -36.1814
25.5463 | -35.6649 | -36.1873
25.5524 | -35.6653 | -36.1929
25.5585 | -35.6657 | -36.1982
25.5646 | -35.6661 | -36.2039
25.5707 | -35.6665 | -36.2097
25.5769 | -35.6668 | -36.2155
25.583 | -35.6672 | -36.2213
25.5891 | -35.6676 | -36.2268
25.5952 | -35.668 | -36.2325
25.6013 | -35.6684 | -36.2381
25.6074 | -35.6689 | -36.2435
25.6135 | -35.6693 | -36.2491
25.6196 | -35.6696 | -36.2548
25.6257 | -35.67 | -36.2603
25.6318 | -35.6705 | -36.266
25.6379 | -35.6709 | -36.2718
25.644 | -35.6712 | -36.2775
25.6501 | -35.6716 | -36.2832
25.6562 | -35.6721 | -36.2887
25.6623 | -35.6725 | -36.2943
25.6684 | -35.6729 | -36.3
25.6745 | -35.6733 | -36.3058
25.6806 | -35.6737 | -36.3112
25.6867 | -35.6741 | -36.3166
25.6928 | -35.6745 | -36.3223
25.6989 | -35.6748 | -36.3278
25.705 | -35.6752 | -36.3333
25.7111 | -35.6756 | -36.3391
25.7173 | -35.676 | -36.3445
25.7234 | -35.6764 | -36.3502
25.7295 | -35.6768 | -36.3559
25.7356 | -35.6772 | -36.3616
25.7417 | -35.6776 | -36.3673
25.7478 | -35.678 | -36.3728
25.7539 | -35.6784 | -36.3783
25.76 | -35.6788 | -36.3839
25.7661 | -35.6792 | -36.3894
25.7722 | -35.6796 | -36.395
25.7783 | -35.68 | -36.4008
25.7844 | -35.6804 | -36.4063
25.7905 | -35.6807 | -36.4121
25.7966 | -35.6811 | -36.4176
25.8027 | -35.6815 | -36.4235
25.8088 | -35.6819 | -36.4291
25.8149 | -35.6823 | -36.4349
25.821 | -35.6826 | -36.44
25.8271 | -35.683 | -36.4456
25.8332 | -35.6834 | -36.4513
25.8393 | -35.6838 | -36.4569
25.8454 | -35.6842 | -36.4628
25.8515 | -35.6846 | -36.4686
25.8576 | -35.685 | -36.4742
25.8638 | -35.6854 | -36.4798
25.8699 | -35.6858 | -36.4854
25.876 | -35.6862 | -36.4908
25.8821 | -35.6865 | -36.4965
25.8882 | -35.6869 | -36.5015
25.8943 | -35.6873 | -36.5072
25.9004 | -35.6877 | -36.5129
25.9065 | -35.6881 | -36.5184
25.9126 | -35.6885 | -36.524
25.9187 | -35.6888 | -36.5298
25.9248 | -35.6892 | -36.5353
25.9309 | -35.6896 | -36.5409
25.937 | -35.69 | -36.5465
25.9431 | -35.6904 | -36.552
25.9492 | -35.6907 | -36.5576
25.9553 | -35.6911 | -36.5631
25.9614 | -35.6915 | -36.5688
25.9675 | -35.6919 | -36.5745
25.9736 | -35.6923 | -36.5801
25.9797 | -35.6927 | -36.5857
25.9858 | -35.693 | -36.5913
25.9919 | -35.6934 | -36.5968
25.998 | -35.6938 | -36.6021
26.0042 | -35.6942 | -36.6077
26.0103 | -35.6946 | -36.6133
26.0164 | -35.6949 | -36.6191
26.0225 | -35.6954 | -36.6247
26.0286 | -35.6957 | -36.6302
26.0347 | -35.6961 | -36.6354
26.0408 | -35.6965 | -36.6413
26.0469 | -35.6969 | -36.6469
26.053 | -35.6973 | -36.6527
26.0591 | -35.6977 | -36.6583
26.0652 | -35.6981 | -36.6636
26.0713 | -35.6985 | -36.6692
26.0774 | -35.6989 | -36.6746
26.0835 | -35.6993 | -36.6802
26.0896 | -35.6996 | -36.6859
26.0957 | -35.7001 | -36.6915
26.1018 | -35.7005 | -36.6969
26.1079 | -35.7009 | -36.7027
26.114 | -35.7013 | -36.7084
26.1201 | -35.7017 | -36.714
26.1262 | -35.702 | -36.7199
26.1323 | -35.7024 | -36.7255
26.1384 | -35.7029 | -36.7308
26.1445 | -35.7032 | -36.7361
26.1507 | -35.7036 | -36.7417
26.1568 | -35.704 | -36.7474
26.1629 | -35.7044 | -36.7531
26.169 | -35.7048 | -36.7587
26.1751 | -35.7051 | -36.7642
26.1812 | -35.7055 | -36.7699
26.1873 | -35.7059 | -36.7755
26.1934 | -35.7063 | -36.7809
26.1995 | -35.7067 | -36.7863
26.2056 | -35.7071 | -36.7918
26.2117 | -35.7075 | -36.7973
26.2178 | -35.7079 | -36.8027
26.2239 | -35.7083 | -36.8083
26.23 | -35.7087 | -36.8139
26.2361 | -35.7091 | -36.8195
26.2422 | -35.7095 | -36.8248
26.2483 | -35.7099 | -36.8304
26.2544 | -35.7102 | -36.8359
26.2605 | -35.7106 | -36.8414
26.2666 | -35.711 | -36.8471
26.2727 | -35.7114 | -36.8525
26.2788 | -35.7118 | -36.8579
26.2849 | -35.7121 | -36.8633
26.2911 | -35.7126 | -36.8688
26.2972 | -35.7129 | -36.8741
26.3033 | -35.7133 | -36.8796
26.3094 | -35.7137 | -36.8852
26.3155 | -35.714 | -36.8907
26.3216 | -35.7144 | -36.8964
26.3277 | -35.7148 | -36.9019
26.3338 | -35.7151 | -36.9075
26.3399 | -35.7155 | -36.9128
26.346 | -35.7159 | -36.9181
26.3521 | -35.7163 | -36.9234
26.3582 | -35.7167 | -36.9288
26.3643 | -35.7171 | -36.9344
26.3704 | -35.7175 | -36.9399
26.3765 | -35.7179 | -36.9455
26.3826 | -35.7182 | -36.9506
26.3887 | -35.7186 | -36.9561
26.3948 | -35.719 | -36.9617
26.4009 | -35.7193 | -36.9673
26.407 | -35.7197 | -36.9727
26.4131 | -35.7201 | -36.9781
26.4192 | -35.7205 | -36.9833
26.4253 | -35.7209 | -36.9888
26.4314 | -35.7212 | -36.9941
26.4376 | -35.7216 | -36.9998
26.4437 | -35.722 | -37.0054
26.4498 | -35.7223 | -37.0109
26.4559 | -35.7227 | -37.016
26.462 | -35.7231 | -37.0218
26.4681 | -35.7234 | -37.0274
26.4742 | -35.7238 | -37.0328
26.4803 | -35.7242 | -37.0383
26.4864 | -35.7246 | -37.0435
26.4925 | -35.725 | -37.0489
26.4986 | -35.7254 | -37.0543
26.5047 | -35.7257 | -37.0598
26.5108 | -35.7261 | -37.0653
26.5169 | -35.7265 | -37.0709
26.523 | -35.7269 | -37.0758
26.5291 | -35.7273 | -37.0811
26.5352 | -35.7276 | -37.0864
26.5413 | -35.728 | -37.0919
26.5474 | -35.7284 | -37.0975
26.5535 | -35.7288 | -37.1032
26.5596 | -35.7292 | -37.1086
26.5657 | -35.7296 | -37.1139
26.5718 | -35.73 | -37.1195
26.578 | -35.7303 | -37.1251
26.5841 | -35.7307 | -37.1306
26.5902 | -35.7311 | -37.1364
26.5963 | -35.7315 | -37.1414
26.6024 | -35.7319 | -37.147
26.6085 | -35.7322 | -37.1523
26.6146 | -35.7326 | -37.1579
26.6207 | -35.733 | -37.1633
26.6268 | -35.7333 | -37.1688
26.6329 | -35.7337 | -37.1738
26.639 | -35.734 | -37.1793
26.6451 | -35.7344 | -37.1849
26.6512 | -35.7348 | -37.19
26.6573 | -35.7352 | -37.1954
26.6634 | -35.7356 | -37.2007
26.6695 | -35.736 | -37.2061
26.6756 | -35.7364 | -37.2114
26.6817 | -35.7367 | -37.2168
26.6878 | -35.7372 | -37.2223
26.6939 | -35.7375 | -37.2278
26.7 | -35.7379 | -37.2329
26.7061 | -35.7383 | -37.2382
26.7122 | -35.7387 | -37.2438
26.7183 | -35.7391 | -37.2492
26.7245 | -35.7395 | -37.2548
26.7306 | -35.7399 | -37.2601
26.7367 | -35.7403 | -37.2655
26.7428 | -35.7407 | -37.2707
26.7489 | -35.741 | -37.276
26.755 | -35.7414 | -37.2813
26.7611 | -35.7418 | -37.2868
26.7672 | -35.7422 | -37.2918
26.7733 | -35.7426 | -37.297
26.7794 | -35.7429 | -37.3026
26.7855 | -35.7433 | -37.3079
26.7916 | -35.7437 | -37.3131
26.7977 | -35.7441 | -37.318
26.8038 | -35.7445 | -37.3231
26.8099 | -35.7449 | -37.3285
26.816 | -35.7452 | -37.3338
26.8221 | -35.7456 | -37.3392
26.8282 | -35.746 | -37.3442
26.8343 | -35.7463 | -37.3495
26.8404 | -35.7467 | -37.3546
26.8465 | -35.747 | -37.3599
26.8526 | -35.7474 | -37.3652
26.8587 | -35.7478 | -37.3708
26.8649 | -35.7482 | -37.3765
26.871 | -35.7486 | -37.3816
26.8771 | -35.749 | -37.3867
26.8832 | -35.7493 | -37.3919
26.8893 | -35.7497 | -37.3972
26.8954 | -35.75 | -37.4028
26.9015 | -35.7504 | -37.4082
26.9076 | -35.7507 | -37.4135
26.9137 | -35.7512 | -37.4187
26.9198 | -35.7516 | -37.4241
26.9259 | -35.752 | -37.4298
26.932 | -35.7523 | -37.4354
26.9381 | -35.7527 | -37.4407
26.9442 | -35.753 | -37.4456
26.9503 | -35.7534 | -37.451
26.9564 | -35.7537 | -37.4559
26.9625 | -35.7542 | -37.4613
26.9686 | -35.7545 | -37.4666
26.9747 | -35.7549 | -37.4722
26.9808 | -35.7553 | -37.477
26.9869 | -35.7556 | -37.4822
26.993 | -35.756 | -37.4876
26.9991 | -35.7564 | -37.4929
27.0052 | -35.7567 | -37.4984
27.0114 | -35.7571 | -37.5039
27.0175 | -35.7575 | -37.5089
27.0236 | -35.7579 | -37.5143
27.0297 | -35.7583 | -37.5195
27.0358 | -35.7586 | -37.5249
27.0419 | -35.759 | -37.5302
27.048 | -35.7593 | -37.5354
27.0541 | -35.7597 | -37.5401
27.0602 | -35.7601 | -37.545
27.0663 | -35.7604 | -37.5503
27.0724 | -35.7608 | -37.5554
27.0785 | -35.7611 | -37.5606
27.0846 | -35.7615 | -37.566
27.0907 | -35.762 | -37.5716
27.0968 | -35.7623 | -37.5768
27.1029 | -35.7627 | -37.5821
27.109 | -35.7631 | -37.5872
27.1151 | -35.7634 | -37.5925
27.1212 | -35.7638 | -37.5976
27.1273 | -35.7642 | -37.6028
27.1334 | -35.7645 | -37.6082
27.1395 | -35.7649 | -37.6138
27.1456 | -35.7653 | -37.6192
27.1518 | -35.7656 | -37.6244
27.1579 | -35.766 | -37.6295
27.164 | -35.7663 | -37.6349
27.1701 | -35.7667 | -37.6401
27.1762 | -35.767 | -37.6454
27.1823 | -35.7674 | -37.6507
27.1884 | -35.7678 | -37.6557
27.1945 | -35.7682 | -37.6604
27.2006 | -35.7686 | -37.6657
27.2067 | -35.769 | -37.6711
27.2128 | -35.7693 | -37.6764
27.2189 | -35.7697 | -37.6819
27.225 | -35.7701 | -37.6873
27.2311 | -35.7705 | -37.6924
27.2372 | -35.7709 | -37.6977
27.2433 | -35.7712 | -37.7028
27.2494 | -35.7715 | -37.7082
27.2555 | -35.7719 | -37.7138
27.2616 | -35.7722 | -37.7188
27.2677 | -35.7726 | -37.7243
27.2738 | -35.773 | -37.7293
27.2799 | -35.7733 | -37.7344
27.286 | -35.7737 | -37.7395
27.2921 | -35.774 | -37.7447
27.2983 | -35.7743 | -37.7499
27.3044 | -35.7747 | -37.7552
27.3105 | -35.775 | -37.7609
27.3166 | -35.7754 | -37.7659
27.3227 | -35.7758 | -37.7715
27.3288 | -35.7762 | -37.7765
27.3349 | -35.7765 | -37.7811
27.341 | -35.7769 | -37.7865
27.3471 | -35.7773 | -37.7918
27.3532 | -35.7776 | -37.7968
27.3593 | -35.7779 | -37.8021
27.3654 | -35.7783 | -37.8073
27.3715 | -35.7787 | -37.8123
27.3776 | -35.7791 | -37.8174
27.3837 | -35.7795 | -37.8225
27.3898 | -35.7798 | -37.8279
27.3959 | -35.7802 | -37.8329
27.402 | -35.7805 | -37.8377
27.4081 | -35.7808 | -37.8431
27.4142 | -35.7811 | -37.8482
27.4203 | -35.7815 | -37.8537
27.4264 | -35.7819 | -37.8588
27.4325 | -35.7822 | -37.8639
27.4387 | -35.7826 | -37.869
27.4448 | -35.783 | -37.8742
27.4509 | -35.7834 | -37.8796
27.457 | -35.7838 | -37.8848
27.4631 | -35.7841 | -37.8902
27.4692 | -35.7845 | -37.8956
27.4753 | -35.7848 | -37.9004
27.4814 | -35.7852 | -37.9057
27.4875 | -35.7855 | -37.9109
27.4936 | -35.7859 | -37.916
27.4997 | -35.7863 | -37.9214
27.5058 | -35.7866 | -37.9266
27.5119 | -35.787 | -37.9319
27.518 | -35.7874 | -37.9371
27.5241 | -35.7877 | -37.9421
27.5302 | -35.7881 | -37.9479
27.5363 | -35.7884 | -37.9528
27.5424 | -35.7888 | -37.9575
27.5485 | -35.7892 | -37.9629
27.5546 | -35.7896 | -37.968
27.5607 | -35.79 | -37.9734
27.5668 | -35.7903 | -37.9787
27.5729 | -35.7906 | -37.9842
27.5791 | -35.7909 | -37.9894
27.5852 | -35.7912 | -37.9947
27.5913 | -35.7916 | -37.9998
27.5974 | -35.792 | -38.005
27.6035 | -35.7924 | -38.0102
27.6096 | -35.7927 | -38.0153
27.6157 | -35.793 | -38.0208
27.6218 | -35.7934 | -38.0258
27.6279 | -35.7938 | -38.0308
27.634 | -35.7941 | -38.0363
27.6401 | -35.7945 | -38.0412
27.6462 | -35.7949 | -38.0463
27.6523 | -35.7952 | -38.0509
27.6584 | -35.7956 | -38.056
27.6645 | -35.7959 | -38.0609
27.6706 | -35.7963 | -38.0662
27.6767 | -35.7967 | -38.071
27.6828 | -35.7971 | -38.0761
27.6889 | -35.7974 | -38.0813
27.695 | -35.7978 | -38.0865
27.7011 | -35.7981 | -38.0919
27.7072 | -35.7985 | -38.0968
27.7133 | -35.7989 | -38.1019
27.7194 | -35.7992 | -38.1068
27.7256 | -35.7996 | -38.1121
27.7317 | -35.7999 | -38.1169
27.7378 | -35.8004 | -38.1223
27.7439 | -35.8007 | -38.1276
27.75 | -35.8011 | -38.1329
27.7561 | -35.8014 | -38.1379
27.7622 | -35.8018 | -38.1432
27.7683 | -35.8021 | -38.1482
27.7744 | -35.8025 | -38.1534
27.7805 | -35.8028 | -38.1586
27.7866 | -35.8032 | -38.1638
27.7927 | -35.8035 | -38.1683
27.7988 | -35.8038 | -38.1736
27.8049 | -35.8042 | -38.1787
27.811 | -35.8045 | -38.1839
27.8171 | -35.8049 | -38.189
27.8232 | -35.8052 | -38.1939
27.8293 | -35.8056 | -38.199
27.8354 | -35.806 | -38.2041
27.8415 | -35.8064 | -38.2093
27.8476 | -35.8068 | -38.2146
27.8537 | -35.8071 | -38.2197
27.8598 | -35.8075 | -38.2245
27.866 | -35.8078 | -38.2298
27.8721 | -35.8081 | -38.2346
27.8782 | -35.8085 | -38.2397
27.8843 | -35.8089 | -38.2448
27.8904 | -35.8093 | -38.25
27.8965 | -35.8097 | -38.2552
27.9026 | -35.8101 | -38.2602
27.9087 | -35.8104 | -38.2658
27.9148 | -35.8108 | -38.2711
27.9209 | -35.8111 | -38.276
27.927 | -35.8114 | -38.2811
27.9331 | -35.8117 | -38.2863
27.9392 | -35.8121 | -38.2914
27.9453 | -35.8124 | -38.2966
27.9514 | -35.8127 | -38.3015
27.9575 | -35.8131 | -38.3069
27.9636 | -35.8134 | -38.3118
27.9697 | -35.8138 | -38.3163
27.9758 | -35.8141 | -38.3217
27.9819 | -35.8145 | -38.3267
27.988 | -35.8148 | -38.3316
27.9941 | -35.8152 | -38.3364
28.0002 | -35.8155 | -38.341
28.0063 | -35.8158 | -38.3461
28.0125 | -35.8161 | -38.3517
28.0186 | -35.8165 | -38.3568
28.0247 | -35.8168 | -38.3621
28.0308 | -35.8172 | -38.367
28.0369 | -35.8175 | -38.372
28.043 | -35.8178 | -38.3773
28.0491 | -35.8182 | -38.3824
28.0552 | -35.8185 | -38.3876
28.0613 | -35.8188 | -38.3924
28.0674 | -35.8191 | -38.3976
28.0735 | -35.8195 | -38.4023

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5302 | -42.5866
21.4908 | -35.5326 | -35.3334
16.4948 | -34.9016 | -27.8896
12.6602 | -34.5388 | -21.5734
9.7171 | -34.3284 | -16.6293
7.4582 | -34.1812 | -12.6337
5.7244 | -34.0713 | -9.3866
4.3936 | -33.9968 | -6.7762
3.3722 | -33.9763 | -4.7165
2.5883 | -33.9827 | -3.1522
1.9866 | -33.9731 | -1.9078
1.5248 | -33.9708 | -0.90806
1.1703 | -33.9899 | -0.25817
1 | -33.9917 | 0.068468
0.89822 | -33.9795 | 0.34397
0.68941 | -33.9719 | 0.92461
0.52919 | -33.964 | 1.3789
0.40612 | -33.967 | 1.987
0.31171 | -33.9686 | 2.4401
0.23928 | -33.9918 | 3.1653
0.18358 | -33.9902 | 3.5848
0.14091 | -33.9902 | 4.4669
0.10819 | -34.0301 | 5.6242
0.083017 | -34.0459 | 7.1047
0.063705 | -34.0776 | 8.9859
0.048923 | -34.1459 | 11.4893
0.037479 | -34.2493 | 14.6896
0.028801 | -34.4388 | 18.7153
0.022078 | -34.7435 | 24.0179
0.017023 | -35.2234 | 30.411
0.013018 | -36.0253 | 38.0921
0.010014 | -37.1298 | 45.996

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003759 | -7.8628
0.007517 | -26.7935
0.011276 | -32.8604
0.015034 | -35.9781
0.018793 | -39.3144
0.022551 | -42.9984
0.02631 | -44.3001
0.030068 | -44.7546
0.033827 | -46.8246
0.037585 | -50.1481
0.041344 | -51.4429
0.045102 | -51.9625
0.048861 | -52.3828
0.052619 | -53.7869
0.056378 | -54.1735
0.060136 | -55.7987
0.063895 | -56.803
0.067653 | -56.6816
0.071412 | -57.1639
0.07517 | -57.8834
0.078929 | -57.9371
0.082687 | -59.1078
0.086446 | -60.3494
0.090204 | -60.2808
0.093963 | -61.0836
0.097721 | -61.7195
0.10148 | -62.03
0.10524 | -63.2512
0.109 | -63.5176
0.11275 | -63.5809
0.11651 | -63.4354
0.12027 | -63.9375
0.12403 | -64.276
0.12779 | -65.2083
0.13155 | -65.2732
0.13531 | -64.9879
0.13906 | -63.9652
0.14282 | -65.4789
0.14658 | -66.7974
0.15034 | -68.0561
0.1541 | -67.223
0.15786 | -66.9598
0.16162 | -67.1527
0.16537 | -67.9424
0.16913 | -68.5143
0.17289 | -68.2734
0.17665 | -68.0251
0.18041 | -68.7767
0.18417 | -69.6759
0.18793 | -69.4958
0.19168 | -70.4499
0.19544 | -70.1478
0.1992 | -69.3458
0.20296 | -69.5036
0.20672 | -69.3549
0.21048 | -70.3396
0.21424 | -71.7024
0.21799 | -72.5991
0.22175 | -72.1047
0.22551 | -71.7343
0.22927 | -72.0008
0.23303 | -72.2158
0.23679 | -73.2674
0.24055 | -73.6087
0.2443 | -73.8656
0.24806 | -73.6042
0.25182 | -73.3736
0.25558 | -73.2641
0.25934 | -73.0816
0.2631 | -73.9224
0.26686 | -74.9792
0.27061 | -74.8097
0.27437 | -74.235
0.27813 | -73.9554
0.28189 | -74.3501
0.28565 | -74.4189
0.28941 | -74.6615
0.29316 | -75.7159
0.29692 | -75.6678
0.30068 | -75.946
0.30444 | -76.5531
0.3082 | -75.8786
0.31196 | -75.2531
0.31572 | -75.9153
0.31947 | -75.6688
0.32323 | -75.7023
0.32699 | -75.6327
0.33075 | -76.335
0.33451 | -77.3951
0.33827 | -76.5233
0.34203 | -75.8511
0.34578 | -76.6286
0.34954 | -76.9805
0.3533 | -77.5896
0.35706 | -77.6501
0.36082 | -76.1735
0.36458 | -76.66
0.36834 | -77.471
0.37209 | -77.2645
0.37585 | -77.836
0.37961 | -78.2554
0.38337 | -77.7262
0.38713 | -77.4934
0.39089 | -77.9097
0.39464 | -77.6779
0.3984 | -77.8531
0.40216 | -77.9683
0.40592 | -78.2706
0.40968 | -78.4463
0.41344 | -77.8692
0.41719 | -77.7634
0.42095 | -77.3315
0.42471 | -78.2889
0.42847 | -78.79
0.43223 | -77.9774
0.43599 | -77.5199
0.43975 | -78.9632
0.4435 | -78.8184
0.44726 | -79.4328
0.45102 | -79.5256
0.45478 | -78.1508
0.45854 | -77.9903
0.4623 | -79.7235
0.46606 | -79.4838
0.46981 | -78.735
0.47357 | -78.7221
0.47733 | -79.08
0.48109 | -80.1707
0.48485 | -80.8146
0.48861 | -81.1519
0.49237 | -81.1996
0.49612 | -79.9963
0.49988 | -79.6123
0.50364 | -79.5279
0.5074 | -79.6727
0.51116 | -79.9112
0.51492 | -80.7402
0.51867 | -80.4123
0.52243 | -80.9601
0.52619 | -80.6666
0.52995 | -80.6549
0.53371 | -80.8085
0.53747 | -80.2986
0.54123 | -80.5278
0.54499 | -81.3086
0.54874 | -81.8265
0.5525 | -81.5375
0.55626 | -81.4863
0.56002 | -81.5485
0.56378 | -81.2398
0.56754 | -80.8624
0.57129 | -80.9902
0.57505 | -80.9994
0.57881 | -80.9628
0.58257 | -81.4342
0.58633 | -82.1656
0.59009 | -82.1833
0.59384 | -81.4369
0.5976 | -81.2929
0.60136 | -81.6817
0.60512 | -81.8109
0.60888 | -81.8912
0.61264 | -82.1697
0.6164 | -82.4594
0.62016 | -82.3715
0.62391 | -82.2815
0.62767 | -82.6071
0.63143 | -83.0957
0.63519 | -83.1671
0.63895 | -82.868
0.64271 | -82.8461
0.64646 | -83.051
0.65022 | -83.0717
0.65398 | -82.5749
0.65774 | -82.0832
0.6615 | -82.1062
0.66526 | -82.3972
0.66902 | -82.7909
0.67277 | -83.4263
0.67653 | -83.7983
0.68029 | -83.8816
0.68405 | -84.2468
0.68781 | -84.2044
0.69157 | -83.8398
0.69532 | -83.5111
0.69908 | -83.3947
0.70284 | -83.59
0.7066 | -83.9134
0.71036 | -83.5935
0.71412 | -83.4782
0.71788 | -83.4524
0.72164 | -83.346
0.72539 | -83.1667
0.72915 | -83.221
0.73291 | -83.9749
0.73667 | -84.9546
0.74043 | -84.7885
0.74419 | -84.0317
0.74794 | -83.9388
0.7517 | -84.2337
0.75546 | -84.7701
0.75922 | -85.2397
0.76298 | -85.0555
0.76674 | -84.7091
0.7705 | -84.7104
0.77425 | -85.2499
0.77801 | -85.525
0.78177 | -85.2772
0.78553 | -84.5688
0.78929 | -83.8541
0.79305 | -83.6908
0.7968 | -83.9727
0.80056 | -84.4489
0.80432 | -85.1528
0.80808 | -85.7194
0.81184 | -85.4643
0.8156 | -85.2924
0.81936 | -85.3605
0.82312 | -85.4267
0.82687 | -85.2502
0.83063 | -84.9637
0.83439 | -85.3001
0.83815 | -85.6536
0.84191 | -85.5149
0.84567 | -85.2429
0.84942 | -84.9387
0.85318 | -84.9935
0.85694 | -85.3131
0.8607 | -85.8656
0.86446 | -86.3293
0.86822 | -86.241
0.87198 | -86.0879
0.87573 | -86.3214
0.87949 | -86.2963
0.88325 | -86.0456
0.88701 | -86.0908
0.89077 | -86.1753
0.89453 | -86.362
0.89829 | -86.4895
0.90204 | -86.6053
0.9058 | -86.6505
0.90956 | -86.1051
0.91332 | -85.3721
0.91708 | -85.221
0.92084 | -85.3042
0.92459 | -85.3458
0.92835 | -85.6223
0.93211 | -86.0352
0.93587 | -86.2637
0.93963 | -86.1586
0.94339 | -86.1468
0.94715 | -86.1119
0.9509 | -85.9044
0.95466 | -85.8617
0.95842 | -86.1436
0.96218 | -86.4522
0.96594 | -86.6173
0.9697 | -86.7421
0.97345 | -86.7211
0.97721 | -86.5282
0.98097 | -86.7916
0.98473 | -87.3107
0.98849 | -87.4404
0.99225 | -87.5216
0.99601 | -87.6327
0.99977 | -87.6679
1.0035 | -87.6753
1.0073 | -87.4456
1.011 | -86.9721
1.0148 | -86.7676
1.0186 | -86.4247
1.0223 | -86.2015
1.0261 | -86.6529
1.0298 | -87.2842
1.0336 | -87.7784
1.0373 | -88.3768
1.0411 | -88.6118
1.0449 | -88.2207
1.0486 | -87.7899
1.0524 | -87.3199
1.0561 | -87.1703
1.0599 | -87.3969
1.0637 | -87.5728
1.0674 | -87.5136
1.0712 | -86.9484
1.0749 | -86.4922
1.0787 | -86.7513
1.0825 | -87.0843
1.0862 | -87.3652
1.09 | -87.6257
1.0937 | -87.7928
1.0975 | -87.7648
1.1012 | -87.6408
1.105 | -87.5982
1.1088 | -87.74
1.1125 | -88.1058
1.1163 | -88.5878
1.12 | -88.8689
1.1238 | -88.6523
1.1276 | -88.1481
1.1313 | -87.8623
1.1351 | -87.9758
1.1388 | -88.2961
1.1426 | -88.4937
1.1463 | -88.4345
1.1501 | -88.2779
1.1539 | -88.1492
1.1576 | -88.0818
1.1614 | -88.1409
1.1651 | -88.282
1.1689 | -88.3528
1.1727 | -88.2991
1.1764 | -88.2825
1.1802 | -88.4203
1.1839 | -88.5692
1.1877 | -88.6429
1.1914 | -88.8249
1.1952 | -89.0926
1.199 | -89.2287
1.2027 | -89.305
1.2065 | -89.414
1.2102 | -89.4867
1.214 | -89.4953
1.2178 | -89.2754
1.2215 | -88.8377
1.2253 | -88.5409
1.229 | -88.6253
1.2328 | -88.8668
1.2366 | -88.8574
1.2403 | -88.6853
1.2441 | -88.7099
1.2478 | -89.0575
1.2516 | -89.5701
1.2553 | -89.861
1.2591 | -89.7364
1.2629 | -89.4187
1.2666 | -89.1545
1.2704 | -88.9949
1.2741 | -88.8485
1.2779 | -88.7456
1.2817 | -88.7837
1.2854 | -88.8893
1.2892 | -89.0563
1.2929 | -89.425
1.2967 | -89.8935
1.3004 | -90.2384
1.3042 | -90.3891
1.308 | -90.3295
1.3117 | -90.1284
1.3155 | -90.0151
1.3192 | -90.0475
1.323 | -90.0788
1.3268 | -90.052
1.3305 | -89.9645
1.3343 | -89.847
1.338 | -89.7482
1.3418 | -89.6108
1.3455 | -89.4559
1.3493 | -89.3689
1.3531 | -89.4032
1.3568 | -89.5979
1.3606 | -89.832
1.3643 | -89.8388
1.3681 | -89.6509
1.3719 | -89.5992
1.3756 | -89.7453
1.3794 | -89.8543
1.3831 | -89.8943
1.3869 | -90.0313
1.3907 | -90.2304
1.3944 | -90.3093
1.3982 | -90.2617
1.4019 | -90.2787
1.4057 | -90.5329
1.4094 | -90.9979
1.4132 | -91.4403
1.417 | -91.5824
1.4207 | -91.4073
1.4245 | -91.1354
1.4282 | -90.9146
1.432 | -90.795
1.4358 | -90.732
1.4395 | -90.6018
1.4433 | -90.4178
1.447 | -90.2723
1.4508 | -90.2331
1.4545 | -90.2206
1.4583 | -90.0976
1.4621 | -90.0007
1.4658 | -90.1134
1.4696 | -90.3478
1.4733 | -90.5329
1.4771 | -90.5931
1.4809 | -90.56
1.4846 | -90.4909
1.4884 | -90.5202
1.4921 | -90.7298
1.4959 | -91.0798
1.4996 | -91.4503
1.5034 | -91.6322
1.5072 | -91.5336
1.5109 | -91.2975
1.5147 | -91.0163
1.5184 | -90.8464
1.5222 | -90.9547
1.526 | -91.1685
1.5297 | -91.1799
1.5335 | -91.0861
1.5372 | -91.0724
1.541 | -91.1131
1.5448 | -91.2573
1.5485 | -91.5824
1.5523 | -91.9016
1.556 | -91.9246
1.5598 | -91.6592
1.5635 | -91.3668
1.5673 | -91.2164
1.5711 | -91.2374
1.5748 | -91.4188
1.5786 | -91.6212
1.5823 | -91.7367
1.5861 | -91.8659
1.5899 | -92.0276
1.5936 | -92.1134
1.5974 | -91.9875
1.6011 | -91.6885
1.6049 | -91.472
1.6086 | -91.4535
1.6124 | -91.4835
1.6162 | -91.4119
1.6199 | -91.4254
1.6237 | -91.697
1.6274 | -91.813
1.6312 | -91.7378
1.635 | -91.5117
1.6387 | -91.299
1.6425 | -91.2261
1.6462 | -91.3696
1.65 | -91.6865
1.6537 | -91.9821
1.6575 | -92.0964
1.6613 | -92.0631
1.665 | -91.9921
1.6688 | -91.914
1.6725 | -91.8385
1.6763 | -91.7989
1.6801 | -91.8175
1.6838 | -91.9161
1.6876 | -92.1113
1.6913 | -92.3445
1.6951 | -92.4536
1.6988 | -92.3049
1.7026 | -91.9734
1.7064 | -91.6697
1.7101 | -91.4991
1.7139 | -91.4526
1.7176 | -91.4906
1.7214 | -91.6145
1.7252 | -91.8174
1.7289 | -92.0407
1.7327 | -92.1737
1.7364 | -92.1781
1.7402 | -92.1216
1.744 | -92.0874
1.7477 | -92.1237
1.7515 | -92.2167
1.7552 | -92.2847
1.759 | -92.2951
1.7627 | -92.2842
1.7665 | -92.2934
1.7703 | -92.328
1.774 | -92.4165
1.7778 | -92.584
1.7815 | -92.7901
1.7853 | -92.9999
1.7891 | -93.1955
1.7928 | -93.3427
1.7966 | -93.4018
1.8003 | -93.314
1.8041 | -93.1244
1.8078 | -92.957
1.8116 | -92.9094
1.8154 | -92.9754
1.8191 | -93.0644
1.8229 | -93.1203
1.8266 | -93.1111
1.8304 | -92.9937
1.8342 | -92.8185
1.8379 | -92.6734
1.8417 | -92.6343
1.8454 | -92.7087
1.8492 | -92.8459
1.8529 | -92.9686
1.8567 | -93.0088
1.8605 | -92.9234
1.8642 | -92.7158
1.868 | -92.4785
1.8717 | -92.3613
1.8755 | -92.3892
1.8793 | -92.4748
1.883 | -92.5604
1.8868 | -92.629
1.8905 | -92.6781
1.8943 | -92.7066
1.8981 | -92.7311
1.9018 | -92.7697
1.9056 | -92.7919
1.9093 | -92.769
1.9131 | -92.6699
1.9168 | -92.5348
1.9206 | -92.4344
1.9244 | -92.4236
1.9281 | -92.5295
1.9319 | -92.6519
1.9356 | -92.6603
1.9394 | -92.5812
1.9432 | -92.563
1.9469 | -92.6892
1.9507 | -92.8908
1.9544 | -93.0911
1.9582 | -93.2343
1.9619 | -93.2265
1.9657 | -93.0016
1.9695 | -92.6294
1.9732 | -92.3231
1.977 | -92.2478
1.9807 | -92.3928
1.9845 | -92.6461
1.9883 | -92.934
1.992 | -93.1962
1.9958 | -93.307
1.9995 | -93.1556
2.0033 | -92.8358
2.007 | -92.5851
2.0108 | -92.5754
2.0146 | -92.832
2.0183 | -93.2527
2.0221 | -93.6679
2.0258 | -93.9805
2.0296 | -94.1296
2.0334 | -94.0778
2.0371 | -93.8901
2.0409 | -93.714
2.0446 | -93.6628
2.0484 | -93.7068
2.0521 | -93.695
2.0559 | -93.5521
2.0597 | -93.4049
2.0634 | -93.4515
2.0672 | -93.6796
2.0709 | -93.869
2.0747 | -93.8357
2.0785 | -93.629
2.0822 | -93.3953
2.086 | -93.2292
2.0897 | -93.1427
2.0935 | -93.1263
2.0973 | -93.1719
2.101 | -93.2582
2.1048 | -93.3416
2.1085 | -93.3737
2.1123 | -93.3291
2.116 | -93.2498
2.1198 | -93.2011
2.1236 | -93.2339
2.1273 | -93.3487
2.1311 | -93.4942
2.1348 | -93.5846
2.1386 | -93.5669
2.1424 | -93.5089
2.1461 | -93.4888
2.1499 | -93.5174
2.1536 | -93.5638
2.1574 | -93.5706
2.1611 | -93.4978
2.1649 | -93.3494
2.1687 | -93.1543
2.1724 | -92.9877
2.1762 | -92.853
2.1799 | -92.7692
2.1837 | -92.7603
2.1875 | -92.8375
2.1912 | -93.0181
2.195 | -93.3018
2.1987 | -93.6462
2.2025 | -93.9736
2.2062 | -94.1962
2.21 | -94.2948
2.2138 | -94.2427
2.2175 | -94.064
2.2213 | -93.8252
2.225 | -93.6144
2.2288 | -93.5172
2.2326 | -93.5575
2.2363 | -93.7094
2.2401 | -93.883
2.2438 | -93.9834
2.2476 | -93.9657
2.2514 | -93.8515
2.2551 | -93.7254
2.2589 | -93.6451
2.2626 | -93.6249
2.2664 | -93.6879
2.2701 | -93.8325
2.2739 | -94.0414
2.2777 | -94.2622
2.2814 | -94.4328
2.2852 | -94.4782
2.2889 | -94.3631
2.2927 | -94.145
2.2965 | -93.9244
2.3002 | -93.7807
2.304 | -93.7463
2.3077 | -93.8136
2.3115 | -93.9814
2.3152 | -94.1998
2.319 | -94.3764
2.3228 | -94.4144
2.3265 | -94.2947
2.3303 | -94.094
2.334 | -93.8981
2.3378 | -93.7696
2.3416 | -93.7392
2.3453 | -93.8026
2.3491 | -93.9391
2.3528 | -94.117
2.3566 | -94.2891
2.3603 | -94.4069
2.3641 | -94.4509
2.3679 | -94.4224
2.3716 | -94.3303
2.3754 | -94.2108
2.3791 | -94.1104
2.3829 | -94.0774
2.3867 | -94.1333
2.3904 | -94.277
2.3942 | -94.4604
2.3979 | -94.6425
2.4017 | -94.7862
2.4055 | -94.8842
2.4092 | -94.9368
2.413 | -94.9214
2.4167 | -94.8219
2.4205 | -94.672
2.4242 | -94.5378
2.428 | -94.4872
2.4318 | -94.5364
2.4355 | -94.6418
2.4393 | -94.7267
2.443 | -94.7136
2.4468 | -94.5491
2.4506 | -94.2692
2.4543 | -93.9694
2.4581 | -93.7572
2.4618 | -93.7145
2.4656 | -93.8526
2.4693 | -94.1111
2.4731 | -94.3897
2.4769 | -94.6
2.4806 | -94.6829
2.4844 | -94.6639
2.4881 | -94.5889
2.4919 | -94.4991
2.4957 | -94.4271
2.4994 | -94.4052
2.5032 | -94.4583
2.5069 | -94.5961
2.5107 | -94.8075
2.5144 | -95.0861
2.5182 | -95.4024
2.522 | -95.7082
2.5257 | -95.9165
2.5295 | -95.9829
2.5332 | -95.9242
2.537 | -95.8015
2.5408 | -95.6562
2.5445 | -95.5075
2.5483 | -95.4132
2.552 | -95.4088
2.5558 | -95.5067
2.5595 | -95.6962
2.5633 | -95.9083
2.5671 | -96.0504
2.5708 | -96.0611
2.5746 | -95.961
2.5783 | -95.8115
2.5821 | -95.6776
2.5859 | -95.598
2.5896 | -95.5902
2.5934 | -95.6733
2.5971 | -95.8158
2.6009 | -95.9603
2.6047 | -96.0515
2.6084 | -96.0776
2.6122 | -96.0522
2.6159 | -96.0124
2.6197 | -95.9654
2.6234 | -95.9064
2.6272 | -95.7918
2.631 | -95.6025
2.6347 | -95.3979
2.6385 | -95.2321
2.6422 | -95.1219
2.646 | -95.0823
2.6498 | -95.0738
2.6535 | -95.0593
2.6573 | -95.0279
2.661 | -95.0234
2.6648 | -95.0955
2.6685 | -95.2823
2.6723 | -95.5593
2.6761 | -95.8503
2.6798 | -96.118
2.6836 | -96.337
2.6873 | -96.4684
2.6911 | -96.513
2.6949 | -96.4871
2.6986 | -96.4308
2.7024 | -96.402
2.7061 | -96.3931
2.7099 | -96.3417
2.7136 | -96.2775
2.7174 | -96.1778
2.7212 | -96.0577
2.7249 | -95.9253
2.7287 | -95.7936
2.7324 | -95.688
2.7362 | -95.6276
2.74 | -95.6262
2.7437 | -95.6746
2.7475 | -95.7465
2.7512 | -95.8366
2.755 | -95.9236
2.7588 | -96.0045
2.7625 | -96.0764
2.7663 | -96.1342
2.77 | -96.1447
2.7738 | -96.0987
2.7775 | -96.0004
2.7813 | -95.8757
2.7851 | -95.7449
2.7888 | -95.621
2.7926 | -95.5328
2.7963 | -95.4961
2.8001 | -95.5258
2.8039 | -95.6293
2.8076 | -95.777
2.8114 | -95.9268
2.8151 | -96.0607
2.8189 | -96.1635
2.8226 | -96.225
2.8264 | -96.2486
2.8302 | -96.2396
2.8339 | -96.2144
2.8377 | -96.1864
2.8414 | -96.1751
2.8452 | -96.1942
2.849 | -96.2342
2.8527 | -96.2878
2.8565 | -96.3244
2.8602 | -96.3167
2.864 | -96.2424
2.8677 | -96.1033
2.8715 | -95.9376
2.8753 | -95.7905
2.879 | -95.6782
2.8828 | -95.6018
2.8865 | -95.5461
2.8903 | -95.5013
2.8941 | -95.4583
2.8978 | -95.4414
2.9016 | -95.439
2.9053 | -95.4394
2.9091 | -95.4417
2.9128 | -95.4629
2.9166 | -95.5328
2.9204 | -95.6698
2.9241 | -95.8703
2.9279 | -96.0896
2.9316 | -96.278
2.9354 | -96.4059
2.9392 | -96.4684
2.9429 | -96.5125
2.9467 | -96.5431
2.9504 | -96.6013
2.9542 | -96.6923
2.958 | -96.792
2.9617 | -96.8955
2.9655 | -96.9683
2.9692 | -96.9968
2.973 | -96.9895
2.9767 | -96.9764
2.9805 | -96.9617
2.9843 | -96.9609
2.988 | -96.9964
2.9918 | -97.0608
2.9955 | -97.1328
2.9993 | -97.1722
3.0031 | -97.1678
3.0068 | -97.1112
3.0106 | -97.0218
3.0143 | -96.9445
3.0181 | -96.906
3.0218 | -96.9336
3.0256 | -97.0428
3.0294 | -97.2249
3.0331 | -97.4559
3.0369 | -97.6633
3.0406 | -97.773
3.0444 | -97.7336
3.0482 | -97.5704
3.0519 | -97.3555
3.0557 | -97.1516
3.0594 | -97.0149
3.0632 | -96.968
3.0669 | -96.9984
3.0707 | -97.0838
3.0745 | -97.193
3.0782 | -97.2945
3.082 | -97.3566
3.0857 | -97.3828
3.0895 | -97.3773
3.0933 | -97.3389
3.097 | -97.2563
3.1008 | -97.152
3.1045 | -97.084
3.1083 | -97.0628
3.1121 | -97.1072
3.1158 | -97.2209
3.1196 | -97.3588
3.1233 | -97.4838
3.1271 | -97.5587
3.1308 | -97.5818
3.1346 | -97.5653
3.1384 | -97.5076
3.1421 | -97.4333
3.1459 | -97.3391
3.1496 | -97.2454
3.1534 | -97.1803
3.1572 | -97.1804
3.1609 | -97.2479
3.1647 | -97.3542
3.1684 | -97.4601
3.1722 | -97.5056
3.1759 | -97.48
3.1797 | -97.3849
3.1835 | -97.2603
3.1872 | -97.154
3.191 | -97.072
3.1947 | -97.0232
3.1985 | -97.0055
3.2023 | -97.0233
3.206 | -97.0742
3.2098 | -97.1529
3.2135 | -97.2425
3.2173 | -97.2997
3.221 | -97.3113
3.2248 | -97.2651
3.2286 | -97.1981
3.2323 | -97.1432
3.2361 | -97.1129
3.2398 | -97.1129
3.2436 | -97.1203
3.2474 | -97.1176
3.2511 | -97.0767
3.2549 | -97.0423
3.2586 | -96.9978
3.2624 | -96.9508
3.2662 | -96.9039
3.2699 | -96.8691
3.2737 | -96.8487
3.2774 | -96.8662
3.2812 | -96.9161
3.2849 | -96.9931
3.2887 | -97.0823
3.2925 | -97.1602
3.2962 | -97.2243
3.3 | -97.262
3.3037 | -97.2885
3.3075 | -97.3121
3.3113 | -97.3403
3.315 | -97.3826
3.3188 | -97.4477
3.3225 | -97.5232
3.3263 | -97.5991
3.33 | -97.6632
3.3338 | -97.6911
3.3376 | -97.6872
3.3413 | -97.6703
3.3451 | -97.6587
3.3488 | -97.666
3.3526 | -97.6977
3.3564 | -97.7519
3.3601 | -97.8175
3.3639 | -97.8741
3.3676 | -97.9163
3.3714 | -97.932
3.3751 | -97.9209
3.3789 | -97.9126
3.3827 | -97.9141
3.3864 | -97.9181
3.3902 | -97.9295
3.3939 | -97.9369
3.3977 | -97.9266
3.4015 | -97.8979
3.4052 | -97.8607
3.409 | -97.8203
3.4127 | -97.7819
3.4165 | -97.7511
3.4202 | -97.7502
3.424 | -97.7802
3.4278 | -97.846
3.4315 | -97.9397
3.4353 | -98.053
3.439 | -98.1643
3.4428 | -98.2503
3.4466 | -98.2945
3.4503 | -98.2862
3.4541 | -98.2208
3.4578 | -98.1213
3.4616 | -97.9996
3.4654 | -97.8748
3.4691 | -97.7699
3.4729 | -97.6931
3.4766 | -97.6576
3.4804 | -97.6658
3.4841 | -97.725
3.4879 | -97.812
3.4917 | -97.8924
3.4954 | -97.9404
3.4992 | -97.9501
3.5029 | -97.9189
3.5067 | -97.8721
3.5105 | -97.8413
3.5142 | -97.8508
3.518 | -97.9081
3.5217 | -98.0169
3.5255 | -98.1437
3.5292 | -98.2505
3.533 | -98.3052
3.5368 | -98.2764
3.5405 | -98.1638
3.5443 | -98.0154
3.548 | -97.8681
3.5518 | -97.7483
3.5556 | -97.6694
3.5593 | -97.6499
3.5631 | -97.6913
3.5668 | -97.7718
3.5706 | -97.8647
3.5743 | -97.9522
3.5781 | -98.0052
3.5819 | -98.0305
3.5856 | -98.0314
3.5894 | -98.0165
3.5931 | -98.0071
3.5969 | -98.0183
3.6007 | -98.0484
3.6044 | -98.0977
3.6082 | -98.1635
3.6119 | -98.2304
3.6157 | -98.2681
3.6195 | -98.2579
3.6232 | -98.211
3.627 | -98.1419
3.6307 | -98.0795
3.6345 | -98.0389
3.6382 | -98.0191
3.642 | -98.0176
3.6458 | -98.0326
3.6495 | -98.0569
3.6533 | -98.0871
3.657 | -98.1389
3.6608 | -98.2146
3.6646 | -98.3104
3.6683 | -98.4012
3.6721 | -98.4765
3.6758 | -98.5243
3.6796 | -98.5301
3.6833 | -98.4968
3.6871 | -98.4314
3.6909 | -98.3703
3.6946 | -98.3318
3.6984 | -98.3075
3.7021 | -98.2863
3.7059 | -98.272
3.7097 | -98.2643
3.7134 | -98.2626
3.7172 | -98.285
3.7209 | -98.3479
3.7247 | -98.4645
3.7284 | -98.6177
3.7322 | -98.7963
3.736 | -98.9721
3.7397 | -99.1265
3.7435 | -99.2353
3.7472 | -99.2747
3.751 | -99.2533
3.7548 | -99.2042
3.7585 | -99.1421
3.7623 | -99.0702
3.766 | -98.9917
3.7698 | -98.9161
3.7735 | -98.8509
3.7773 | -98.8012
3.7811 | -98.7799
3.7848 | -98.8
3.7886 | -98.8472
3.7923 | -98.9099
3.7961 | -98.9583
3.7999 | -98.9933
3.8036 | -99.0148
3.8074 | -99.0198
3.8111 | -99.012
3.8149 | -98.9837
3.8187 | -98.9391
3.8224 | -98.8842
3.8262 | -98.8239
3.8299 | -98.7641
3.8337 | -98.7076
3.8374 | -98.6666
3.8412 | -98.6436
3.845 | -98.6238
3.8487 | -98.6119
3.8525 | -98.6044
3.8562 | -98.5874
3.86 | -98.5714
3.8638 | -98.5571
3.8675 | -98.5569
3.8713 | -98.5813
3.875 | -98.6287
3.8788 | -98.697
3.8825 | -98.7811
3.8863 | -98.8648
3.8901 | -98.9308
3.8938 | -98.9747
3.8976 | -98.9837
3.9013 | -98.9536
3.9051 | -98.901
3.9089 | -98.8472
3.9126 | -98.8066
3.9164 | -98.7914
3.9201 | -98.802
3.9239 | -98.8253
3.9276 | -98.8662
3.9314 | -98.9085
3.9352 | -98.9336
3.9389 | -98.9461
3.9427 | -98.9479
3.9464 | -98.927
3.9502 | -98.9015
3.954 | -98.8842
3.9577 | -98.8838
3.9615 | -98.9129
3.9652 | -98.9709
3.969 | -99.0426
3.9728 | -99.1101
3.9765 | -99.1451
3.9803 | -99.1434
3.984 | -99.1082
3.9878 | -99.0465
3.9915 | -98.9609
3.9953 | -98.8752
3.9991 | -98.8146
4.0028 | -98.7801
4.0066 | -98.7858
4.0103 | -98.8264
4.0141 | -98.8803
4.0179 | -98.9398
4.0216 | -98.9982
4.0254 | -99.0495
4.0291 | -99.0867
4.0329 | -99.1129
4.0366 | -99.1301
4.0404 | -99.1414
4.0442 | -99.1427
4.0479 | -99.1384
4.0517 | -99.1313
4.0554 | -99.1242
4.0592 | -99.1183
4.063 | -99.1057
4.0667 | -99.0921
4.0705 | -99.0632
4.0742 | -99.033
4.078 | -99.006
4.0817 | -98.9722
4.0855 | -98.9304
4.0893 | -98.9007
4.093 | -98.896
4.0968 | -98.9237
4.1005 | -98.9792
4.1043 | -99.0562
4.1081 | -99.1375
4.1118 | -99.2149
4.1156 | -99.2682
4.1193 | -99.2916
4.1231 | -99.2814
4.1269 | -99.2505
4.1306 | -99.2011
4.1344 | -99.1536
4.1381 | -99.123
4.1419 | -99.1156
4.1456 | -99.1435
4.1494 | -99.1877
4.1532 | -99.2317
4.1569 | -99.2718
4.1607 | -99.2961
4.1644 | -99.3095
4.1682 | -99.3054
4.172 | -99.2929
4.1757 | -99.2708
4.1795 | -99.2549
4.1832 | -99.2561
4.187 | -99.2775
4.1907 | -99.3082
4.1945 | -99.3543
4.1983 | -99.4077
4.202 | -99.4521
4.2058 | -99.4919
4.2095 | -99.5155
4.2133 | -99.5299
4.2171 | -99.5234
4.2208 | -99.5018
4.2246 | -99.4702
4.2283 | -99.4216
4.2321 | -99.3694
4.2358 | -99.3018
4.2396 | -99.2299
4.2434 | -99.1682
4.2471 | -99.1066
4.2509 | -99.0533
4.2546 | -99.0186
4.2584 | -99.0086
4.2622 | -99.0315
4.2659 | -99.0923
4.2697 | -99.1831
4.2734 | -99.2962
4.2772 | -99.4192
4.2809 | -99.53
4.2847 | -99.5987
4.2885 | -99.6256
4.2922 | -99.6141
4.296 | -99.5672
4.2997 | -99.4983
4.3035 | -99.4092
4.3073 | -99.3114
4.311 | -99.2324
4.3148 | -99.1859
4.3185 | -99.1813
4.3223 | -99.217
4.3261 | -99.2881
4.3298 | -99.3894
4.3336 | -99.5088
4.3373 | -99.6138
4.3411 | -99.7133
4.3448 | -99.7998
4.3486 | -99.8618
4.3524 | -99.8954
4.3561 | -99.9013
4.3599 | -99.8804
4.3636 | -99.8342
4.3674 | -99.79
4.3712 | -99.7474
4.3749 | -99.705
4.3787 | -99.6708
4.3824 | -99.6421
4.3862 | -99.6284
4.3899 | -99.6319
4.3937 | -99.6488
4.3975 | -99.6742
4.4012 | -99.7108
4.405 | -99.7531
4.4087 | -99.7923
4.4125 | -99.8259
4.4163 | -99.8504
4.42 | -99.8727
4.4238 | -99.8874
4.4275 | -99.8898
4.4313 | -99.8979
4.435 | -99.915
4.4388 | -99.9367
4.4426 | -99.9841
4.4463 | -100.0426
4.4501 | -100.094
4.4538 | -100.1321
4.4576 | -100.136
4.4614 | -100.0995
4.4651 | -100.0247
4.4689 | -99.9178
4.4726 | -99.7835
4.4764 | -99.6384
4.4802 | -99.5077
4.4839 | -99.4031
4.4877 | -99.3382
4.4914 | -99.3134
4.4952 | -99.3384
4.4989 | -99.4094
4.5027 | -99.5025
4.5065 | -99.6096
4.5102 | -99.7205
4.514 | -99.8095
4.5177 | -99.8855
4.5215 | -99.9452
4.5253 | -99.9868
4.529 | -100.0185
4.5328 | -100.0434
4.5365 | -100.0677
4.5403 | -100.0869
4.544 | -100.1085
4.5478 | -100.1248
4.5516 | -100.1318
4.5553 | -100.135
4.5591 | -100.1326
4.5628 | -100.1344
4.5666 | -100.1277
4.5704 | -100.1196
4.5741 | -100.1246
4.5779 | -100.1304
4.5816 | -100.1307
4.5854 | -100.1239
4.5891 | -100.1023
4.5929 | -100.071
4.5967 | -100.0254
4.6004 | -99.9685
4.6042 | -99.9097
4.6079 | -99.8403
4.6117 | -99.7728
4.6155 | -99.7168
4.6192 | -99.688
4.623 | -99.6943
4.6267 | -99.738
4.6305 | -99.8269
4.6342 | -99.9526
4.638 | -100.111
4.6418 | -100.2748
4.6455 | -100.4223
4.6493 | -100.5634
4.653 | -100.6658
4.6568 | -100.7308
4.6606 | -100.7462
4.6643 | -100.7003
4.6681 | -100.6018
4.6718 | -100.4627
4.6756 | -100.2915
4.6794 | -100.1182
4.6831 | -99.9537
4.6869 | -99.8177
4.6906 | -99.7286
4.6944 | -99.6923
4.6981 | -99.714
4.7019 | -99.7819
4.7057 | -99.8868
4.7094 | -100.0157
4.7132 | -100.1527
4.7169 | -100.2799
4.7207 | -100.3597
4.7245 | -100.4159
4.7282 | -100.4373
4.732 | -100.4413
4.7357 | -100.4463
4.7395 | -100.4515
4.7432 | -100.4575
4.747 | -100.4634
4.7508 | -100.4693
4.7545 | -100.4843
4.7583 | -100.4994
4.762 | -100.5108
4.7658 | -100.5196
4.7696 | -100.5212
4.7733 | -100.5248
4.7771 | -100.5216
4.7808 | -100.5131
4.7846 | -100.5203
4.7883 | -100.5371
4.7921 | -100.5785
4.7959 | -100.6377
4.7996 | -100.7174
4.8034 | -100.8066
4.8071 | -100.8997
4.8109 | -100.9889
4.8147 | -101.0555
4.8184 | -101.0955
4.8222 | -101.0936
4.8259 | -101.043
4.8297 | -100.9632
4.8335 | -100.8639
4.8372 | -100.7554
4.841 | -100.6572
4.8447 | -100.5747
4.8485 | -100.5206
4.8522 | -100.4864
4.856 | -100.4658
4.8598 | -100.4694
4.8635 | -100.4776
4.8673 | -100.4877
4.871 | -100.4939
4.8748 | -100.4952
4.8786 | -100.4918
4.8823 | -100.4928
4.8861 | -100.488
4.8898 | -100.48
4.8936 | -100.4613
4.8973 | -100.4349
4.9011 | -100.4071
4.9049 | -100.3791
4.9086 | -100.3561
4.9124 | -100.3445
4.9161 | -100.3424
4.9199 | -100.3433
4.9237 | -100.3514
4.9274 | -100.3725
4.9312 | -100.4063
4.9349 | -100.4495
4.9387 | -100.5113
4.9424 | -100.5835
4.9462 | -100.6685
4.95 | -100.753
4.9537 | -100.8286
4.9575 | -100.8872
4.9612 | -100.9275
4.965 | -100.9467
4.9688 | -100.9439
4.9725 | -100.9235
4.9763 | -100.8832
4.98 | -100.8343
4.9838 | -100.7778
4.9875 | -100.7241
4.9913 | -100.6726
4.9951 | -100.6252
4.9988 | -100.5904
5.0026 | -100.5585
5.0063 | -100.5343
5.0101 | -100.507
5.0139 | -100.4882
5.0176 | -100.4781
5.0214 | -100.4741
5.0251 | -100.4747
5.0289 | -100.4826
5.0327 | -100.4894
5.0364 | -100.5032
5.0402 | -100.5295
5.0439 | -100.5504
5.0477 | -100.5674
5.0514 | -100.5863
5.0552 | -100.6025
5.059 | -100.6052
5.0627 | -100.5956
5.0665 | -100.5866
5.0702 | -100.5805
5.074 | -100.5786
5.0778 | -100.5885
5.0815 | -100.6099
5.0853 | -100.6369
5.089 | -100.6718
5.0928 | -100.7017
5.0965 | -100.7219
5.1003 | -100.7373
5.1041 | -100.7423
5.1078 | -100.732
5.1116 | -100.7098
5.1153 | -100.6804
5.1191 | -100.6448
5.1229 | -100.6236
5.1266 | -100.6171
5.1304 | -100.6243
5.1341 | -100.6632
5.1379 | -100.7331
5.1416 | -100.8265
5.1454 | -100.9346
5.1492 | -101.0379
5.1529 | -101.1407
5.1567 | -101.2226
5.1604 | -101.2731
5.1642 | -101.2744
5.168 | -101.2293
5.1717 | -101.1443
5.1755 | -101.0324
5.1792 | -100.9065
5.183 | -100.7809
5.1868 | -100.6756
5.1905 | -100.5984
5.1943 | -100.5485
5.198 | -100.5293
5.2018 | -100.5428
5.2055 | -100.5827
5.2093 | -100.6458
5.2131 | -100.7151
5.2168 | -100.7817
5.2206 | -100.8455
5.2243 | -100.8852
5.2281 | -100.9143
5.2319 | -100.917
5.2356 | -100.9067
5.2394 | -100.8975
5.2431 | -100.9038
5.2469 | -100.9203
5.2506 | -100.9437
5.2544 | -100.977
5.2582 | -101.0168
5.2619 | -101.0565
5.2657 | -101.0897
5.2694 | -101.113
5.2732 | -101.132
5.277 | -101.1471
5.2807 | -101.1499
5.2845 | -101.1468
5.2882 | -101.1409
5.292 | -101.1301
5.2957 | -101.1205
5.2995 | -101.104
5.3033 | -101.086
5.307 | -101.067
5.3108 | -101.0542
5.3145 | -101.0559
5.3183 | -101.0695
5.3221 | -101.0866
5.3258 | -101.1059
5.3296 | -101.1366
5.3333 | -101.1784
5.3371 | -101.2174
5.3408 | -101.2563
5.3446 | -101.2882
5.3484 | -101.3124
5.3521 | -101.3254
5.3559 | -101.3274
5.3596 | -101.3383
5.3634 | -101.3434
5.3672 | -101.3479
5.3709 | -101.3594
5.3747 | -101.3661
5.3784 | -101.3773
5.3822 | -101.3928
5.386 | -101.4042
5.3897 | -101.4069
5.3935 | -101.3998
5.3972 | -101.3948
5.401 | -101.3948
5.4047 | -101.3927
5.4085 | -101.3794
5.4123 | -101.3589
5.416 | -101.3433
5.4198 | -101.345
5.4235 | -101.3449
5.4273 | -101.3535
5.4311 | -101.3635
5.4348 | -101.386
5.4386 | -101.4197
5.4423 | -101.4451
5.4461 | -101.4835
5.4498 | -101.5253
5.4536 | -101.5657
5.4574 | -101.6144
5.4611 | -101.6665
5.4649 | -101.7155
5.4686 | -101.7728
5.4724 | -101.8335
5.4762 | -101.888
5.4799 | -101.9393
5.4837 | -101.991
5.4874 | -102.041
5.4912 | -102.0764
5.4949 | -102.0965
5.4987 | -102.1126
5.5025 | -102.1159
5.5062 | -102.1101
5.51 | -102.0911
5.5137 | -102.0656
5.5175 | -102.04
5.5213 | -102.0081
5.525 | -101.9698
5.5288 | -101.9303
5.5325 | -101.8947
5.5363 | -101.8606
5.5401 | -101.8333
5.5438 | -101.8102
5.5476 | -101.7875
5.5513 | -101.7724
5.5551 | -101.7553
5.5588 | -101.7436
5.5626 | -101.7359
5.5664 | -101.7278
5.5701 | -101.7109
5.5739 | -101.7002
5.5776 | -101.7001
5.5814 | -101.7023
5.5852 | -101.7034
5.5889 | -101.7016
5.5927 | -101.7045
5.5964 | -101.7113
5.6002 | -101.7128
5.6039 | -101.7086
5.6077 | -101.7065
5.6115 | -101.6987
5.6152 | -101.6875
5.619 | -101.6661
5.6227 | -101.6398
5.6265 | -101.6094
5.6303 | -101.5696
5.634 | -101.5341
5.6378 | -101.5109
5.6415 | -101.5013
5.6453 | -101.5041
5.649 | -101.5252
5.6528 | -101.5621
5.6566 | -101.6177
5.6603 | -101.6737
5.6641 | -101.7299
5.6678 | -101.7872
5.6716 | -101.8319
5.6754 | -101.8514
5.6791 | -101.8343
5.6829 | -101.7871
5.6866 | -101.7169
5.6904 | -101.6205
5.6942 | -101.5032
5.6979 | -101.396
5.7017 | -101.297
5.7054 | -101.2169
5.7092 | -101.1637
5.7129 | -101.1468
5.7167 | -101.1625
5.7205 | -101.2011
5.7242 | -101.2635
5.728 | -101.3453
5.7317 | -101.4443
5.7355 | -101.5526
5.7393 | -101.6625
5.743 | -101.77
5.7468 | -101.8671
5.7505 | -101.947
5.7543 | -102.0054
5.758 | -102.051
5.7618 | -102.0887
5.7656 | -102.1101
5.7693 | -102.1158
5.7731 | -102.1093
5.7768 | -102.1072
5.7806 | -102.1072
5.7844 | -102.1089
5.7881 | -102.1246
5.7919 | -102.1493
5.7956 | -102.1915
5.7994 | -102.2405
5.8031 | -102.3033
5.8069 | -102.3697
5.8107 | -102.4288
5.8144 | -102.4821
5.8182 | -102.5171
5.8219 | -102.5395
5.8257 | -102.5475
5.8295 | -102.5494
5.8332 | -102.5555
5.837 | -102.5564
5.8407 | -102.5595
5.8445 | -102.5588
5.8483 | -102.5508
5.852 | -102.5296
5.8558 | -102.4948
5.8595 | -102.4446
5.8633 | -102.3733
5.867 | -102.2829
5.8708 | -102.189
5.8746 | -102.1014
5.8783 | -102.0265
5.8821 | -101.9575
5.8858 | -101.9064
5.8896 | -101.8743
5.8934 | -101.8651
5.8971 | -101.8688
5.9009 | -101.8877
5.9046 | -101.9268
5.9084 | -101.9757
5.9121 | -102.037
5.9159 | -102.0984
5.9197 | -102.1537
5.9234 | -102.1982
5.9272 | -102.2313
5.9309 | -102.2495
5.9347 | -102.2599
5.9385 | -102.2591
5.9422 | -102.2469
5.946 | -102.221
5.9497 | -102.1957
5.9535 | -102.1742
5.9572 | -102.1433
5.961 | -102.1083
5.9648 | -102.0643
5.9685 | -102.0158
5.9723 | -101.9755
5.976 | -101.9291
5.9798 | -101.8792
5.9836 | -101.8268
5.9873 | -101.7803
5.9911 | -101.7416
5.9948 | -101.7056
5.9986 | -101.6842
6.0023 | -101.6767
6.0061 | -101.6804
6.0099 | -101.6953
6.0136 | -101.7295
6.0174 | -101.7695
6.0211 | -101.8186
6.0249 | -101.871
6.0287 | -101.9226
6.0324 | -101.9711
6.0362 | -102.0219
6.0399 | -102.0558
6.0437 | -102.0724
6.0475 | -102.0832
6.0512 | -102.0893
6.055 | -102.0918
6.0587 | -102.0906
6.0625 | -102.08
6.0662 | -102.0631
6.07 | -102.0485
6.0738 | -102.037
6.0775 | -102.0287
6.0813 | -102.0259
6.085 | -102.0354
6.0888 | -102.053
6.0926 | -102.0711
6.0963 | -102.0956
6.1001 | -102.1283
6.1038 | -102.1631
6.1076 | -102.1876
6.1113 | -102.1975
6.1151 | -102.1963
6.1189 | -102.1794
6.1226 | -102.1636
6.1264 | -102.1378
6.1301 | -102.1004
6.1339 | -102.0665
6.1377 | -102.0296
6.1414 | -101.9905
6.1452 | -101.9526
6.1489 | -101.9278
6.1527 | -101.9123
6.1564 | -101.9116
6.1602 | -101.9263
6.164 | -101.9577
6.1677 | -102.0053
6.1715 | -102.0671
6.1752 | -102.1367
6.179 | -102.2024
6.1828 | -102.2694
6.1865 | -102.3246
6.1903 | -102.3541
6.194 | -102.3608
6.1978 | -102.346
6.2016 | -102.3179
6.2053 | -102.2805
6.2091 | -102.2373
6.2128 | -102.1954
6.2166 | -102.1497
6.2203 | -102.1114
6.2241 | -102.0885
6.2279 | -102.079
6.2316 | -102.0822
6.2354 | -102.0935
6.2391 | -102.1136
6.2429 | -102.1389
6.2467 | -102.1748
6.2504 | -102.2041
6.2542 | -102.2261
6.2579 | -102.2424
6.2617 | -102.2465
6.2654 | -102.2365
6.2692 | -102.2198
6.273 | -102.1966
6.2767 | -102.1709
6.2805 | -102.143
6.2842 | -102.1154
6.288 | -102.0931
6.2918 | -102.0731
6.2955 | -102.0532
6.2993 | -102.0394
6.303 | -102.0271
6.3068 | -102.0221
6.3105 | -102.0233
6.3143 | -102.0303
6.3181 | -102.041
6.3218 | -102.0561
6.3256 | -102.0845
6.3293 | -102.1141
6.3331 | -102.1411
6.3369 | -102.1777
6.3406 | -102.2127
6.3444 | -102.2446
6.3481 | -102.2622
6.3519 | -102.2632
6.3556 | -102.2542
6.3594 | -102.2301
6.3632 | -102.1948
6.3669 | -102.1505
6.3707 | -102.0986
6.3744 | -102.0513
6.3782 | -102.0017
6.382 | -101.9605
6.3857 | -101.9305
6.3895 | -101.912
6.3932 | -101.9052
6.397 | -101.9149
6.4008 | -101.9386
6.4045 | -101.9756
6.4083 | -102.0252
6.412 | -102.0826
6.4158 | -102.1528
6.4195 | -102.2281
6.4233 | -102.3131
6.4271 | -102.4008
6.4308 | -102.4804
6.4346 | -102.5543
6.4383 | -102.6245
6.4421 | -102.6883
6.4459 | -102.7382
6.4496 | -102.7687
6.4534 | -102.7872
6.4571 | -102.7822
6.4609 | -102.7589
6.4646 | -102.7267
6.4684 | -102.6878
6.4722 | -102.6444
6.4759 | -102.5893
6.4797 | -102.5391
6.4834 | -102.4921
6.4872 | -102.4553
6.491 | -102.4229
6.4947 | -102.3907
6.4985 | -102.3694
6.5022 | -102.3596
6.506 | -102.3607
6.5097 | -102.3575
6.5135 | -102.3548
6.5173 | -102.3545
6.521 | -102.3557
6.5248 | -102.3598
6.5285 | -102.369
6.5323 | -102.3829
6.5361 | -102.3979
6.5398 | -102.4133
6.5436 | -102.4321
6.5473 | -102.4549
6.5511 | -102.4745
6.5549 | -102.4988
6.5586 | -102.523
6.5624 | -102.5363
6.5661 | -102.5469
6.5699 | -102.5616
6.5736 | -102.5711
6.5774 | -102.574
6.5812 | -102.5722
6.5849 | -102.5575
6.5887 | -102.5387
6.5924 | -102.524
6.5962 | -102.5003
6.6 | -102.4675
6.6037 | -102.4392
6.6075 | -102.4186
6.6112 | -102.3995
6.615 | -102.3858
6.6187 | -102.3788
6.6225 | -102.3794
6.6263 | -102.384
6.63 | -102.3894
6.6338 | -102.3993
6.6375 | -102.4144
6.6413 | -102.4289
6.6451 | -102.4408
6.6488 | -102.4523
6.6526 | -102.4654
6.6563 | -102.4747
6.6601 | -102.4842
6.6638 | -102.4929
6.6676 | -102.4905
6.6714 | -102.4907
6.6751 | -102.4902
6.6789 | -102.4881
6.6826 | -102.4867
6.6864 | -102.4844
6.6902 | -102.4749
6.6939 | -102.4576
6.6977 | -102.4453
6.7014 | -102.4346
6.7052 | -102.4272
6.7089 | -102.4204
6.7127 | -102.4179
6.7165 | -102.4181
6.7202 | -102.4205
6.724 | -102.422
6.7277 | -102.4275
6.7315 | -102.4368
6.7353 | -102.4508
6.739 | -102.4723
6.7428 | -102.498
6.7465 | -102.5263
6.7503 | -102.5575
6.7541 | -102.5905
6.7578 | -102.6249
6.7616 | -102.6569
6.7653 | -102.6804
6.7691 | -102.7056
6.7728 | -102.7182
6.7766 | -102.7368
6.7804 | -102.7531
6.7841 | -102.7641
6.7879 | -102.7751
6.7916 | -102.7895
6.7954 | -102.7966
6.7992 | -102.7901
6.8029 | -102.7852
6.8067 | -102.7731
6.8104 | -102.7592
6.8142 | -102.7489
6.8179 | -102.7434
6.8217 | -102.7423
6.8255 | -102.7465
6.8292 | -102.7519
6.833 | -102.7605
6.8367 | -102.768
6.8405 | -102.765
6.8443 | -102.7636
6.848 | -102.7608
6.8518 | -102.7479
6.8555 | -102.7267
6.8593 | -102.7002
6.863 | -102.6768
6.8668 | -102.6631
6.8706 | -102.6411
6.8743 | -102.6297
6.8781 | -102.6261
6.8818 | -102.6284
6.8856 | -102.6441
6.8894 | -102.6704
6.8931 | -102.702
6.8969 | -102.7396
6.9006 | -102.7787
6.9044 | -102.8147
6.9082 | -102.8508
6.9119 | -102.8815
6.9157 | -102.9017
6.9194 | -102.9151
6.9232 | -102.9181
6.9269 | -102.9057
6.9307 | -102.8786
6.9345 | -102.8406
6.9382 | -102.7962
6.942 | -102.7485
6.9457 | -102.6972
6.9495 | -102.6409
6.9533 | -102.592
6.957 | -102.5553
6.9608 | -102.5268
6.9645 | -102.4996
6.9683 | -102.4771
6.972 | -102.4671
6.9758 | -102.461
6.9796 | -102.4548
6.9833 | -102.4574
6.9871 | -102.4687
6.9908 | -102.4879
6.9946 | -102.5149
6.9984 | -102.556
7.0021 | -102.6027
7.0059 | -102.6482
7.0096 | -102.6939
7.0134 | -102.74
7.0171 | -102.7869
7.0209 | -102.823
7.0247 | -102.8533
7.0284 | -102.8762
7.0322 | -102.8907
7.0359 | -102.8948
7.0397 | -102.895
7.0435 | -102.8942
7.0472 | -102.8859
7.051 | -102.8836
7.0547 | -102.8856
7.0585 | -102.8861
7.0622 | -102.8854
7.066 | -102.8935
7.0698 | -102.909
7.0735 | -102.9282
7.0773 | -102.9532
7.081 | -102.9744
7.0848 | -103.002
7.0886 | -103.0311
7.0923 | -103.0521
7.0961 | -103.0696
7.0998 | -103.0785
7.1036 | -103.0792
7.1074 | -103.0831
7.1111 | -103.0849
7.1149 | -103.0749
7.1186 | -103.06
7.1224 | -103.0376
7.1261 | -103.0112
7.1299 | -102.9838
7.1337 | -102.9555
7.1374 | -102.9262
7.1412 | -102.8981
7.1449 | -102.8684
7.1487 | -102.8379
7.1525 | -102.8199
7.1562 | -102.8078
7.16 | -102.7997
7.1637 | -102.7984
7.1675 | -102.8014
7.1712 | -102.8087
7.175 | -102.8182
7.1788 | -102.8364
7.1825 | -102.8587
7.1863 | -102.8851
7.19 | -102.9202
7.1938 | -102.9506
7.1976 | -102.9809
7.2013 | -103.011
7.2051 | -103.0343
7.2088 | -103.0554
7.2126 | -103.0669
7.2164 | -103.0673
7.2201 | -103.0584
7.2239 | -103.0506
7.2276 | -103.0372
7.2314 | -103.0177
7.2351 | -103.0005
7.2389 | -102.9838
7.2427 | -102.9683
7.2464 | -102.9559
7.2502 | -102.9469
7.2539 | -102.9491
7.2577 | -102.9604
7.2615 | -102.9731
7.2652 | -102.9988
7.269 | -103.0325
7.2727 | -103.0722
7.2765 | -103.1148
7.2802 | -103.1556
7.284 | -103.188
7.2878 | -103.2164
7.2915 | -103.2403
7.2953 | -103.2503
7.299 | -103.2529
7.3028 | -103.2481
7.3066 | -103.2344
7.3103 | -103.2204
7.3141 | -103.2065
7.3178 | -103.1908
7.3216 | -103.1745
7.3253 | -103.1585
7.3291 | -103.1526
7.3329 | -103.1491
7.3366 | -103.1508
7.3404 | -103.1559
7.3441 | -103.1617
7.3479 | -103.1714
7.3517 | -103.1814
7.3554 | -103.1914
7.3592 | -103.199
7.3629 | -103.1992
7.3667 | -103.1987
7.3704 | -103.1908
7.3742 | -103.178
7.378 | -103.1676
7.3817 | -103.1611
7.3855 | -103.1596
7.3892 | -103.1605
7.393 | -103.167
7.3968 | -103.1716
7.4005 | -103.1796
7.4043 | -103.1973
7.408 | -103.2122
7.4118 | -103.2239
7.4156 | -103.2303
7.4193 | -103.2255
7.4231 | -103.2158
7.4268 | -103.2008
7.4306 | -103.1817
7.4343 | -103.165
7.4381 | -103.1474
7.4419 | -103.1353
7.4456 | -103.1316
7.4494 | -103.1391
7.4531 | -103.152
7.4569 | -103.1707
7.4607 | -103.1954
7.4644 | -103.2202
7.4682 | -103.2467
7.4719 | -103.2703
7.4757 | -103.2831
7.4794 | -103.2895
7.4832 | -103.2857
7.487 | -103.2724
7.4907 | -103.2465
7.4945 | -103.2117
7.4982 | -103.1753
7.502 | -103.1307
7.5058 | -103.0867
7.5095 | -103.0409
7.5133 | -102.9973
7.517 | -102.964
7.5208 | -102.9478
7.5245 | -102.9428
7.5283 | -102.9551
7.5321 | -102.9789
7.5358 | -103.0156
7.5396 | -103.0647
7.5433 | -103.1201
7.5471 | -103.1812
7.5509 | -103.2446
7.5546 | -103.3062
7.5584 | -103.3588
7.5621 | -103.4084
7.5659 | -103.4487
7.5697 | -103.476
7.5734 | -103.4925
7.5772 | -103.5034
7.5809 | -103.5034
7.5847 | -103.4972
7.5884 | -103.4766
7.5922 | -103.4659
7.596 | -103.4529
7.5997 | -103.4316
7.6035 | -103.4118
7.6072 | -103.3928
7.611 | -103.3749
7.6148 | -103.3579
7.6185 | -103.3406
7.6223 | -103.3278
7.626 | -103.3087
7.6298 | -103.2832
7.6335 | -103.2588
7.6373 | -103.2396
7.6411 | -103.2168
7.6448 | -103.1947
7.6486 | -103.1693
7.6523 | -103.1441
7.6561 | -103.1177
7.6599 | -103.0937
7.6636 | -103.0701
7.6674 | -103.0461
7.6711 | -103.0294
7.6749 | -103.0193
7.6786 | -103.0137
7.6824 | -103.0228
7.6862 | -103.0432
7.6899 | -103.0683
7.6937 | -103.1053
7.6974 | -103.1507
7.7012 | -103.1956
7.705 | -103.2476
7.7087 | -103.2993
7.7125 | -103.3457
7.7162 | -103.3946
7.72 | -103.4318
7.7237 | -103.46
7.7275 | -103.476
7.7313 | -103.4839
7.735 | -103.4828
7.7388 | -103.4737
7.7425 | -103.4627
7.7463 | -103.4486
7.7501 | -103.4293
7.7538 | -103.409
7.7576 | -103.3967
7.7613 | -103.3865
7.7651 | -103.3861
7.7689 | -103.3899
7.7726 | -103.4025
7.7764 | -103.4205
7.7801 | -103.4478
7.7839 | -103.4826
7.7876 | -103.5211
7.7914 | -103.5571
7.7952 | -103.5938
7.7989 | -103.629
7.8027 | -103.6595
7.8064 | -103.6853
7.8102 | -103.703
7.814 | -103.7128
7.8177 | -103.7154
7.8215 | -103.7061
7.8252 | -103.6924
7.829 | -103.6731
7.8327 | -103.6435
7.8365 | -103.6139
7.8403 | -103.585
7.844 | -103.5559
7.8478 | -103.527
7.8515 | -103.5032
7.8553 | -103.4867
7.8591 | -103.4766
7.8628 | -103.471
7.8666 | -103.4652
7.8703 | -103.4632
7.8741 | -103.4619
7.8778 | -103.4718
7.8816 | -103.4716
7.8854 | -103.47
7.8891 | -103.4733
7.8929 | -103.4716
7.8966 | -103.4612
7.9004 | -103.4443
7.9042 | -103.4251
7.9079 | -103.4069
7.9117 | -103.3801
7.9154 | -103.3489
7.9192 | -103.3195
7.923 | -103.2914
7.9267 | -103.2613
7.9305 | -103.2317
7.9342 | -103.2028
7.938 | -103.1775
7.9417 | -103.1566
7.9455 | -103.1385
7.9493 | -103.1242
7.953 | -103.1131
7.9568 | -103.1073
7.9605 | -103.109
7.9643 | -103.115
7.9681 | -103.1223
7.9718 | -103.1292
7.9756 | -103.1445
7.9793 | -103.1616
7.9831 | -103.1814
7.9868 | -103.1987
7.9906 | -103.2169
7.9944 | -103.2269
7.9981 | -103.2416
8.0019 | -103.252
8.0056 | -103.2538
8.0094 | -103.2576
8.0132 | -103.2628
8.0169 | -103.2742
8.0207 | -103.2808
8.0244 | -103.29
8.0282 | -103.3075
8.0319 | -103.319
8.0357 | -103.3314
8.0395 | -103.3436
8.0432 | -103.3623
8.047 | -103.3858
8.0507 | -103.4085
8.0545 | -103.4299
8.0583 | -103.4516
8.062 | -103.473
8.0658 | -103.494
8.0695 | -103.5156
8.0733 | -103.534
8.077 | -103.5526
8.0808 | -103.572
8.0846 | -103.5911
8.0883 | -103.6136
8.0921 | -103.6253
8.0958 | -103.6398
8.0996 | -103.6568
8.1034 | -103.6714
8.1071 | -103.6826
8.1109 | -103.6967
8.1146 | -103.7087
8.1184 | -103.7197
8.1222 | -103.7271
8.1259 | -103.7312
8.1297 | -103.7351
8.1334 | -103.7405
8.1372 | -103.74
8.1409 | -103.7481
8.1447 | -103.7546
8.1485 | -103.7662
8.1522 | -103.7841
8.156 | -103.7999
8.1597 | -103.8191
8.1635 | -103.8325
8.1673 | -103.8426
8.171 | -103.8557
8.1748 | -103.8633
8.1785 | -103.8643
8.1823 | -103.8718
8.186 | -103.8792
8.1898 | -103.8775
8.1936 | -103.8736
8.1973 | -103.8694
8.2011 | -103.867
8.2048 | -103.8614
8.2086 | -103.8539
8.2124 | -103.8517
8.2161 | -103.8452
8.2199 | -103.8329
8.2236 | -103.8247
8.2274 | -103.822
8.2311 | -103.8147
8.2349 | -103.8008
8.2387 | -103.7857
8.2424 | -103.7751
8.2462 | -103.7617
8.2499 | -103.7481
8.2537 | -103.7306
8.2575 | -103.7101
8.2612 | -103.6915
8.265 | -103.6668
8.2687 | -103.6454
8.2725 | -103.6245
8.2763 | -103.6063
8.28 | -103.5887
8.2838 | -103.5748
8.2875 | -103.5679
8.2913 | -103.5623
8.295 | -103.5615
8.2988 | -103.5729
8.3026 | -103.5859
8.3063 | -103.5976
8.3101 | -103.6172
8.3138 | -103.642
8.3176 | -103.6686
8.3214 | -103.6973
8.3251 | -103.7269
8.3289 | -103.7617
8.3326 | -103.7959
8.3364 | -103.8262
8.3401 | -103.8562
8.3439 | -103.8827
8.3477 | -103.9056
8.3514 | -103.9204
8.3552 | -103.939
8.3589 | -103.9612
8.3627 | -103.98
8.3665 | -103.9966
8.3702 | -104.0181
8.374 | -104.0347
8.3777 | -104.0546
8.3815 | -104.0812
8.3852 | -104.104
8.389 | -104.129
8.3928 | -104.1456
8.3965 | -104.1603
8.4003 | -104.1715
8.404 | -104.1798
8.4078 | -104.1837
8.4116 | -104.1816
8.4153 | -104.1767
8.4191 | -104.1623
8.4228 | -104.14
8.4266 | -104.1269
8.4304 | -104.1074
8.4341 | -104.0878
8.4379 | -104.0677
8.4416 | -104.0462
8.4454 | -104.0297
8.4491 | -104.0153
8.4529 | -103.9997
8.4567 | -103.9865
8.4604 | -103.9776
8.4642 | -103.9744
8.4679 | -103.9746
8.4717 | -103.9731
8.4755 | -103.9764
8.4792 | -103.9759
8.483 | -103.9822
8.4867 | -103.998
8.4905 | -104.0146
8.4942 | -104.0291
8.498 | -104.0467
8.5018 | -104.0621
8.5055 | -104.068
8.5093 | -104.0775
8.513 | -104.0863
8.5168 | -104.0949
8.5206 | -104.0974
8.5243 | -104.0963
8.5281 | -104.0936
8.5318 | -104.0939
8.5356 | -104.09
8.5393 | -104.0837
8.5431 | -104.0833
8.5469 | -104.084
8.5506 | -104.0733
8.5544 | -104.0682
8.5581 | -104.0579
8.5619 | -104.0502
8.5657 | -104.0431
8.5694 | -104.0301
8.5732 | -104.016
8.5769 | -104.0052
8.5807 | -103.9938
8.5844 | -103.988
8.5882 | -103.9797
8.592 | -103.9758
8.5957 | -103.9806
8.5995 | -103.988
8.6032 | -104.0035
8.607 | -104.0188
8.6108 | -104.0286
8.6145 | -104.0413
8.6183 | -104.0599
8.622 | -104.0795
8.6258 | -104.0998
8.6296 | -104.1171
8.6333 | -104.1286
8.6371 | -104.1364
8.6408 | -104.1402
8.6446 | -104.1396
8.6483 | -104.1363
8.6521 | -104.1316
8.6559 | -104.1215
8.6596 | -104.1149
8.6634 | -104.1109
8.6671 | -104.1062
8.6709 | -104.1049
8.6747 | -104.1246
8.6784 | -104.1341
8.6822 | -104.1435
8.6859 | -104.1593
8.6897 | -104.1784
8.6934 | -104.2021
8.6972 | -104.2275
8.701 | -104.2585
8.7047 | -104.2876
8.7085 | -104.317
8.7122 | -104.3447
8.716 | -104.3683
8.7198 | -104.3893
8.7235 | -104.4036
8.7273 | -104.4103
8.731 | -104.4121
8.7348 | -104.4092
8.7385 | -104.4023
8.7423 | -104.3942
8.7461 | -104.3778
8.7498 | -104.3546
8.7536 | -104.334
8.7573 | -104.3147
8.7611 | -104.2917
8.7649 | -104.2701
8.7686 | -104.2557
8.7724 | -104.2425
8.7761 | -104.23
8.7799 | -104.213
8.7836 | -104.2029
8.7874 | -104.2016
8.7912 | -104.203
8.7949 | -104.2025
8.7987 | -104.2017
8.8024 | -104.2047
8.8062 | -104.2031
8.81 | -104.2009
8.8137 | -104.1985
8.8175 | -104.1896
8.8212 | -104.1759
8.825 | -104.1607
8.8288 | -104.1419
8.8325 | -104.1228
8.8363 | -104.1079
8.84 | -104.0952
8.8438 | -104.0841
8.8475 | -104.0812
8.8513 | -104.079
8.8551 | -104.0817
8.8588 | -104.0922
8.8626 | -104.1071
8.8663 | -104.1289
8.8701 | -104.1579
8.8739 | -104.1845
8.8776 | -104.2181
8.8814 | -104.2535
8.8851 | -104.2873
8.8889 | -104.3242
8.8926 | -104.3616
8.8964 | -104.3952
8.9002 | -104.4218
8.9039 | -104.4507
8.9077 | -104.4772
8.9114 | -104.5036
8.9152 | -104.521
8.919 | -104.5352
8.9227 | -104.5459
8.9265 | -104.5617
8.9302 | -104.5799
8.934 | -104.592
8.9377 | -104.6076
8.9415 | -104.6186
8.9453 | -104.6236
8.949 | -104.628
8.9528 | -104.6301
8.9565 | -104.6265
8.9603 | -104.6163
8.9641 | -104.5976
8.9678 | -104.5706
8.9716 | -104.5413
8.9753 | -104.5062
8.9791 | -104.4657
8.9829 | -104.4243
8.9866 | -104.3812
8.9904 | -104.3355
8.9941 | -104.2908
8.9979 | -104.247
9.0016 | -104.2037
9.0054 | -104.1703
9.0092 | -104.1339
9.0129 | -104.1041
9.0167 | -104.077
9.0204 | -104.055
9.0242 | -104.0422
9.028 | -104.0295
9.0317 | -104.0213
9.0355 | -104.0149
9.0392 | -104.0086
9.043 | -104.0058
9.0467 | -104.0043
9.0505 | -104.0027
9.0543 | -104.0052
9.058 | -104.0066
9.0618 | -104.0064
9.0655 | -104.0101
9.0693 | -104.0192
9.0731 | -104.0317
9.0768 | -104.0499
9.0806 | -104.0734
9.0843 | -104.1009
9.0881 | -104.1324
9.0918 | -104.1673
9.0956 | -104.2052
9.0994 | -104.2502
9.1031 | -104.2955
9.1069 | -104.3438
9.1106 | -104.3946
9.1144 | -104.4379
9.1182 | -104.4737
9.1219 | -104.5112
9.1257 | -104.544
9.1294 | -104.5676
9.1332 | -104.5873
9.137 | -104.6002
9.1407 | -104.6169
9.1445 | -104.6253
9.1482 | -104.626
9.152 | -104.6193
9.1557 | -104.6043
9.1595 | -104.5918
9.1633 | -104.5764
9.167 | -104.5559
9.1708 | -104.5366
9.1745 | -104.5169
9.1783 | -104.4959
9.1821 | -104.4745
9.1858 | -104.458
9.1896 | -104.4427
9.1933 | -104.4298
9.1971 | -104.4181
9.2008 | -104.4083
9.2046 | -104.4024
9.2084 | -104.3971
9.2121 | -104.3942
9.2159 | -104.4146
9.2196 | -104.4246
9.2234 | -104.4357
9.2272 | -104.4528
9.2309 | -104.4707
9.2347 | -104.4843
9.2384 | -104.4997
9.2422 | -104.5162
9.2459 | -104.5332
9.2497 | -104.5545
9.2535 | -104.5761
9.2572 | -104.5922
9.261 | -104.6042
9.2647 | -104.6178
9.2685 | -104.6309
9.2723 | -104.6408
9.276 | -104.6443
9.2798 | -104.6483
9.2835 | -104.6528
9.2873 | -104.6603
9.2911 | -104.6632
9.2948 | -104.6598
9.2986 | -104.6567
9.3023 | -104.6516
9.3061 | -104.6449
9.3098 | -104.635
9.3136 | -104.626
9.3174 | -104.6125
9.3211 | -104.5937
9.3249 | -104.5758
9.3286 | -104.5588
9.3324 | -104.5427
9.3362 | -104.5245
9.3399 | -104.505
9.3437 | -104.4836
9.3474 | -104.4627
9.3512 | -104.4359
9.3549 | -104.4095
9.3587 | -104.386
9.3625 | -104.3687
9.3662 | -104.3537
9.37 | -104.3368
9.3737 | -104.3242
9.3775 | -104.313
9.3813 | -104.3019
9.385 | -104.2953
9.3888 | -104.2973
9.3925 | -104.3037
9.3963 | -104.3158
9.4 | -104.3353
9.4038 | -104.3593
9.4076 | -104.3855
9.4113 | -104.4154
9.4151 | -104.4494
9.4188 | -104.4873
9.4226 | -104.5244
9.4264 | -104.5616
9.4301 | -104.6052
9.4339 | -104.6471
9.4376 | -104.6804
9.4414 | -104.7141
9.4451 | -104.7463
9.4489 | -104.7776
9.4527 | -104.8068
9.4564 | -104.8305
9.4602 | -104.8507
9.4639 | -104.8724
9.4677 | -104.8875
9.4715 | -104.8985
9.4752 | -104.9077
9.479 | -104.9151
9.4827 | -104.923
9.4865 | -104.9259
9.4903 | -104.9321
9.494 | -104.9309
9.4978 | -104.9315
9.5015 | -104.9257
9.5053 | -104.9145
9.509 | -104.9062
9.5128 | -104.8954
9.5166 | -104.8802
9.5203 | -104.8687
9.5241 | -104.8517
9.5278 | -104.8271
9.5316 | -104.8004
9.5354 | -104.7718
9.5391 | -104.7423
9.5429 | -104.7112
9.5466 | -104.683
9.5504 | -104.654
9.5541 | -104.621
9.5579 | -104.5893
9.5617 | -104.5628
9.5654 | -104.5396
9.5692 | -104.5196
9.5729 | -104.5064
9.5767 | -104.4994
9.5805 | -104.4922
9.5842 | -104.4893
9.588 | -104.4917
9.5917 | -104.4991
9.5955 | -104.518
9.5992 | -104.5417
9.603 | -104.5644
9.6068 | -104.5933
9.6105 | -104.6261
9.6143 | -104.6608
9.618 | -104.6982
9.6218 | -104.7337
9.6256 | -104.7681
9.6293 | -104.7983
9.6331 | -104.8266
9.6368 | -104.8441
9.6406 | -104.8541
9.6443 | -104.8621
9.6481 | -104.8631
9.6519 | -104.8536
9.6556 | -104.8378
9.6594 | -104.8098
9.6631 | -104.7798
9.6669 | -104.7488
9.6707 | -104.716
9.6744 | -104.6835
9.6782 | -104.6545
9.6819 | -104.6243
9.6857 | -104.5963
9.6895 | -104.575
9.6932 | -104.5647
9.697 | -104.5613
9.7007 | -104.5604
9.7045 | -104.5629
9.7082 | -104.5723
9.712 | -104.5905
9.7158 | -104.6182
9.7195 | -104.6569
9.7233 | -104.7008
9.727 | -104.7469
9.7308 | -104.7966
9.7346 | -104.8482
9.7383 | -104.8979
9.7421 | -104.9483
9.7458 | -105.0003
9.7496 | -105.0492
9.7533 | -105.0929
9.7571 | -105.1098
9.7609 | -105.1401
9.7646 | -105.1652
9.7684 | -105.1798
9.7721 | -105.1873
9.7759 | -105.1932
9.7797 | -105.1978
9.7834 | -105.1967
9.7872 | -105.1864
9.7909 | -105.1697
9.7947 | -105.1578
9.7985 | -105.1475
9.8022 | -105.1372
9.806 | -105.1243
9.8097 | -105.1091
9.8135 | -105.0931
9.8172 | -105.0742
9.821 | -105.0504
9.8248 | -105.0309
9.8285 | -105.0081
9.8323 | -104.9811
9.836 | -104.953
9.8398 | -104.9149
9.8436 | -104.8807
9.8473 | -104.8498
9.8511 | -104.8104
9.8548 | -104.7687
9.8586 | -104.7253
9.8623 | -104.6812
9.8661 | -104.6393
9.8699 | -104.6003
9.8736 | -104.5612
9.8774 | -104.5246
9.8811 | -104.494
9.8849 | -104.4666
9.8887 | -104.4407
9.8924 | -104.4192
9.8962 | -104.4066
9.8999 | -104.4025
9.9037 | -104.403
9.9074 | -104.4121
9.9112 | -104.4248
9.915 | -104.4465
9.9187 | -104.4752
9.9225 | -104.5058
9.9262 | -104.5392
9.93 | -104.5759
9.9338 | -104.6138
9.9375 | -104.6495
9.9413 | -104.6833
9.945 | -104.7169
9.9488 | -104.7447
9.9525 | -104.7712
9.9563 | -104.7941
9.9601 | -104.8096
9.9638 | -104.8216
9.9676 | -104.8298
9.9713 | -104.8322
9.9751 | -104.8261
9.9789 | -104.8187
9.9826 | -104.8037
9.9864 | -104.795
9.9901 | -104.7861
9.9939 | -104.7743
9.9977 | -104.7704
10.0014 | -104.7675
10.0052 | -104.766
10.0089 | -104.768
10.0127 | -104.7727
10.0164 | -104.7794
10.0202 | -104.7935
10.024 | -104.8098
10.0277 | -104.8277
10.0315 | -104.8467
10.0352 | -104.8668
10.039 | -104.8897
10.0428 | -104.9106
10.0465 | -104.9335
10.0503 | -104.9585
10.054 | -104.9804
10.0578 | -105.0005
10.0615 | -105.0168
10.0653 | -105.0275
10.0691 | -105.0364
10.0728 | -105.0441
10.0766 | -105.0535
10.0803 | -105.0578
10.0841 | -105.06
10.0879 | -105.0573
10.0916 | -105.047
10.0954 | -105.0389
10.0991 | -105.0298
10.1029 | -105.0212
10.1066 | -105.013
10.1104 | -105.0021
10.1142 | -104.9897
10.1179 | -104.9764
10.1217 | -104.9697
10.1254 | -104.9632
10.1292 | -104.9623
10.133 | -104.9575
10.1367 | -104.9499
10.1405 | -104.9517
10.1442 | -104.9596
10.148 | -104.9707
10.1517 | -104.9862
10.1555 | -105.005
10.1593 | -105.0292
10.163 | -105.0558
10.1668 | -105.0793
10.1705 | -105.1056
10.1743 | -105.1312
10.1781 | -105.1545
10.1818 | -105.1775
10.1856 | -105.199
10.1893 | -105.2172
10.1931 | -105.2276
10.1969 | -105.2347
10.2006 | -105.2407
10.2044 | -105.2452
10.2081 | -105.2449
10.2119 | -105.2435
10.2156 | -105.2395
10.2194 | -105.2342
10.2232 | -105.2261
10.2269 | -105.2174
10.2307 | -105.2046
10.2344 | -105.1875
10.2382 | -105.1722
10.242 | -105.1595
10.2457 | -105.1432
10.2495 | -105.1246
10.2532 | -105.1079
10.257 | -105.0884
10.2607 | -105.0715
10.2645 | -105.0544
10.2683 | -105.0378
10.272 | -105.0239
10.2758 | -105.0075
10.2795 | -104.9904
10.2833 | -104.9739
10.2871 | -104.9541
10.2908 | -104.9337
10.2946 | -104.9175
10.2983 | -104.8979
10.3021 | -104.8827
10.3058 | -104.8628
10.3096 | -104.8446
10.3134 | -104.8285
10.3171 | -104.8158
10.3209 | -104.8028
10.3246 | -104.7907
10.3284 | -104.7834
10.3322 | -104.7767
10.3359 | -104.7697
10.3397 | -104.7655
10.3434 | -104.7658
10.3472 | -104.7669
10.351 | -104.7711
10.3547 | -104.7793
10.3585 | -104.7888
10.3622 | -104.7994
10.366 | -104.8104
10.3697 | -104.8262
10.3735 | -104.8436
10.3773 | -104.8645
10.381 | -104.8864
10.3848 | -104.9067
10.3885 | -104.924
10.3923 | -104.9328
10.3961 | -104.9391
10.3998 | -104.9446
10.4036 | -104.9504
10.4073 | -104.957
10.4111 | -104.9574
10.4148 | -104.9559
10.4186 | -104.9499
10.4224 | -104.9388
10.4261 | -104.9293
10.4299 | -104.9191
10.4336 | -104.9105
10.4374 | -104.8983
10.4412 | -104.8872
10.4449 | -104.8758
10.4487 | -104.8679
10.4524 | -104.863
10.4562 | -104.8625
10.4599 | -104.869
10.4637 | -104.8751
10.4675 | -104.8839
10.4712 | -104.8951
10.475 | -104.9114
10.4787 | -104.932
10.4825 | -104.9499
10.4863 | -104.9688
10.49 | -104.9904
10.4938 | -105.0132
10.4975 | -105.0363
10.5013 | -105.0588
10.5051 | -105.0778
10.5088 | -105.0968
10.5126 | -105.1165
10.5163 | -105.1331
10.5201 | -105.145
10.5238 | -105.1462
10.5276 | -105.1499
10.5314 | -105.1548
10.5351 | -105.1586
10.5389 | -105.1611
10.5426 | -105.1575
10.5464 | -105.1521
10.5502 | -105.1469
10.5539 | -105.1387
10.5577 | -105.125
10.5614 | -105.1106
10.5652 | -105.0971
10.5689 | -105.0802
10.5727 | -105.0679
10.5765 | -105.058
10.5802 | -105.0493
10.584 | -105.0425
10.5877 | -105.034
10.5915 | -105.0302
10.5953 | -105.028
10.599 | -105.0296
10.6028 | -105.0373
10.6065 | -105.0452
10.6103 | -105.057
10.614 | -105.0709
10.6178 | -105.0878
10.6216 | -105.1055
10.6253 | -105.1253
10.6291 | -105.1453
10.6328 | -105.1679
10.6366 | -105.1932
10.6404 | -105.2169
10.6441 | -105.24
10.6479 | -105.2615
10.6516 | -105.2811
10.6554 | -105.3014
10.6592 | -105.3184
10.6629 | -105.3304
10.6667 | -105.3403
10.6704 | -105.3507
10.6742 | -105.3647
10.6779 | -105.3782
10.6817 | -105.3861
10.6855 | -105.392
10.6892 | -105.3938
10.693 | -105.3925
10.6967 | -105.3946
10.7005 | -105.3944
10.7043 | -105.3914
10.708 | -105.3855
10.7118 | -105.3778
10.7155 | -105.368
10.7193 | -105.3553
10.723 | -105.3419
10.7268 | -105.3251
10.7306 | -105.3086
10.7343 | -105.2906
10.7381 | -105.2717
10.7418 | -105.2541
10.7456 | -105.2377
10.7494 | -105.2199
10.7531 | -105.1999
10.7569 | -105.1791
10.7606 | -105.1576
10.7644 | -105.137
10.7681 | -105.1213
10.7719 | -105.1079
10.7757 | -105.094
10.7794 | -105.0828
10.7832 | -105.0749
10.7869 | -105.0704
10.7907 | -105.071
10.7945 | -105.0712
10.7982 | -105.0711
10.802 | -105.0751
10.8057 | -105.0823
10.8095 | -105.0922
10.8132 | -105.1061
10.817 | -105.1227
10.8208 | -105.1346
10.8245 | -105.1441
10.8283 | -105.1513
10.832 | -105.1571
10.8358 | -105.1604
10.8396 | -105.1576
10.8433 | -105.1445
10.8471 | -105.1427
10.8508 | -105.1424
10.8546 | -105.1387
10.8584 | -105.1325
10.8621 | -105.1215
10.8659 | -105.1122
10.8696 | -105.106
10.8734 | -105.1001
10.8771 | -105.0974
10.8809 | -105.0967
10.8847 | -105.0965
10.8884 | -105.096
10.8922 | -105.0977
10.8959 | -105.1037
10.8997 | -105.1069
10.9035 | -105.1122
10.9072 | -105.1244
10.911 | -105.1398
10.9147 | -105.1576
10.9185 | -105.1794
10.9222 | -105.2018
10.926 | -105.2236
10.9298 | -105.2482
10.9335 | -105.2755
10.9373 | -105.3015
10.941 | -105.326
10.9448 | -105.3494
10.9486 | -105.3735
10.9523 | -105.4016
10.9561 | -105.4275
10.9598 | -105.4531
10.9636 | -105.475
10.9673 | -105.4944
10.9711 | -105.5107
10.9749 | -105.5257
10.9786 | -105.5452
10.9824 | -105.565
10.9861 | -105.5868
10.9899 | -105.6079
10.9937 | -105.6259
10.9974 | -105.6432
11.0012 | -105.6598
11.0049 | -105.6722
11.0087 | -105.6859
11.0124 | -105.6953
11.0162 | -105.7021
11.02 | -105.7094
11.0237 | -105.7179
11.0275 | -105.7247
11.0312 | -105.7238
11.035 | -105.7207
11.0388 | -105.714
11.0425 | -105.7069
11.0463 | -105.6969
11.05 | -105.6855
11.0538 | -105.6722
11.0576 | -105.6545
11.0613 | -105.6397
11.0651 | -105.6244
11.0688 | -105.606
11.0726 | -105.5917
11.0763 | -105.5782
11.0801 | -105.5605
11.0839 | -105.5454
11.0876 | -105.535
11.0914 | -105.5276
11.0951 | -105.5258
11.0989 | -105.5231
11.1027 | -105.5213
11.1064 | -105.5241
11.1102 | -105.53
11.1139 | -105.5353
11.1177 | -105.5424
11.1214 | -105.5457
11.1252 | -105.5525
11.129 | -105.5581
11.1327 | -105.5627
11.1365 | -105.5667
11.1402 | -105.5695
11.144 | -105.5679
11.1478 | -105.5619
11.1515 | -105.5542
11.1553 | -105.5421
11.159 | -105.5303
11.1628 | -105.5159
11.1665 | -105.4977
11.1703 | -105.4762
11.1741 | -105.4575
11.1778 | -105.4397
11.1816 | -105.4179
11.1853 | -105.3964
11.1891 | -105.3754
11.1929 | -105.3592
11.1966 | -105.3442
11.2004 | -105.3321
11.2041 | -105.321
11.2079 | -105.312
11.2117 | -105.3094
11.2154 | -105.3108
11.2192 | -105.3153
11.2229 | -105.3212
11.2267 | -105.3283
11.2304 | -105.3417
11.2342 | -105.358
11.238 | -105.377
11.2417 | -105.4046
11.2455 | -105.4325
11.2492 | -105.459
11.253 | -105.4834
11.2568 | -105.5076
11.2605 | -105.5388
11.2643 | -105.5704
11.268 | -105.5992
11.2718 | -105.6257
11.2755 | -105.6533
11.2793 | -105.6782
11.2831 | -105.7031
11.2868 | -105.7229
11.2906 | -105.7424
11.2943 | -105.759
11.2981 | -105.7699
11.3019 | -105.7769
11.3056 | -105.7814
11.3094 | -105.7871
11.3131 | -105.7899
11.3169 | -105.7936
11.3206 | -105.7926
11.3244 | -105.7863
11.3282 | -105.7782
11.3319 | -105.7685
11.3357 | -105.7515
11.3394 | -105.7303
11.3432 | -105.7169
11.347 | -105.6982
11.3507 | -105.6757
11.3545 | -105.6501
11.3582 | -105.6232
11.362 | -105.5971
11.3658 | -105.571
11.3695 | -105.5462
11.3733 | -105.5154
11.377 | -105.4842
11.3808 | -105.4547
11.3845 | -105.4386
11.3883 | -105.4146
11.3921 | -105.3929
11.3958 | -105.3725
11.3996 | -105.3542
11.4033 | -105.3449
11.4071 | -105.3356
11.4109 | -105.3304
11.4146 | -105.3281
11.4184 | -105.3316
11.4221 | -105.3368
11.4259 | -105.339
11.4296 | -105.3439
11.4334 | -105.3546
11.4372 | -105.3683
11.4409 | -105.3802
11.4447 | -105.3943
11.4484 | -105.4094
11.4522 | -105.4287
11.456 | -105.4452
11.4597 | -105.4612
11.4635 | -105.4827
11.4672 | -105.5003
11.471 | -105.5168
11.4747 | -105.5321
11.4785 | -105.5469
11.4823 | -105.5607
11.486 | -105.5738
11.4898 | -105.5842
11.4935 | -105.5887
11.4973 | -105.5928
11.5011 | -105.5985
11.5048 | -105.5973
11.5086 | -105.5865
11.5123 | -105.577
11.5161 | -105.5673
11.5198 | -105.5595
11.5236 | -105.5494
11.5274 | -105.5349
11.5311 | -105.5169
11.5349 | -105.4976
11.5386 | -105.4818
11.5424 | -105.465
11.5462 | -105.4445
11.5499 | -105.4203
11.5537 | -105.3964
11.5574 | -105.3755
11.5612 | -105.3533
11.565 | -105.3294
11.5687 | -105.307
11.5725 | -105.2829
11.5762 | -105.2621
11.58 | -105.2451
11.5837 | -105.2307
11.5875 | -105.2176
11.5913 | -105.2066
11.595 | -105.2003
11.5988 | -105.1986
11.6025 | -105.198
11.6063 | -105.1945
11.6101 | -105.1988
11.6138 | -105.2069
11.6176 | -105.2148
11.6213 | -105.2207
11.6251 | -105.2333
11.6288 | -105.2481
11.6326 | -105.2599
11.6364 | -105.2742
11.6401 | -105.2908
11.6439 | -105.306
11.6476 | -105.3221
11.6514 | -105.3353
11.6552 | -105.3487
11.6589 | -105.3664
11.6627 | -105.387
11.6664 | -105.4034
11.6702 | -105.417
11.6739 | -105.4353
11.6777 | -105.4542
11.6815 | -105.4712
11.6852 | -105.4862
11.689 | -105.4998
11.6927 | -105.5152
11.6965 | -105.5315
11.7003 | -105.5503
11.704 | -105.5669
11.7078 | -105.5819
11.7115 | -105.5936
11.7153 | -105.6037
11.7191 | -105.612
11.7228 | -105.6176
11.7266 | -105.621
11.7303 | -105.6243
11.7341 | -105.6271
11.7378 | -105.6256
11.7416 | -105.6226
11.7454 | -105.6171
11.7491 | -105.6085
11.7529 | -105.6007
11.7566 | -105.5918
11.7604 | -105.5792
11.7642 | -105.5676
11.7679 | -105.5552
11.7717 | -105.543
11.7754 | -105.5335
11.7792 | -105.5218
11.7829 | -105.5098
11.7867 | -105.5018
11.7905 | -105.4949
11.7942 | -105.4884
11.798 | -105.4871
11.8017 | -105.4888
11.8055 | -105.4931
11.8093 | -105.4973
11.813 | -105.5048
11.8168 | -105.5141
11.8205 | -105.5296
11.8243 | -105.5444
11.828 | -105.5525
11.8318 | -105.561
11.8356 | -105.5747
11.8393 | -105.5898
11.8431 | -105.6047
11.8468 | -105.617
11.8506 | -105.6257
11.8544 | -105.6312
11.8581 | -105.6307
11.8619 | -105.6264
11.8656 | -105.6219
11.8694 | -105.6151
11.8732 | -105.6057
11.8769 | -105.5962
11.8807 | -105.5832
11.8844 | -105.5684
11.8882 | -105.5517
11.8919 | -105.5343
11.8957 | -105.5152
11.8995 | -105.4923
11.9032 | -105.4693
11.907 | -105.4463
11.9107 | -105.4245
11.9145 | -105.4085
11.9183 | -105.3938
11.922 | -105.3806
11.9258 | -105.3834
11.9295 | -105.3762
11.9333 | -105.3721
11.937 | -105.3668
11.9408 | -105.36
11.9446 | -105.3563
11.9483 | -105.3589
11.9521 | -105.3603
11.9558 | -105.3597
11.9596 | -105.3609
11.9634 | -105.3669
11.9671 | -105.3747
11.9709 | -105.3796
11.9746 | -105.3878
11.9784 | -105.3957
11.9821 | -105.4024
11.9859 | -105.4108
11.9897 | -105.4192
11.9934 | -105.4243
11.9972 | -105.4323
12.0009 | -105.4436
12.0047 | -105.4513
12.0085 | -105.455
12.0122 | -105.4571
12.016 | -105.457
12.0197 | -105.459
12.0235 | -105.4617
12.0272 | -105.4672
12.031 | -105.4756
12.0348 | -105.4834
12.0385 | -105.4896
12.0423 | -105.496
12.046 | -105.5044
12.0498 | -105.5121
12.0536 | -105.5253
12.0573 | -105.5357
12.0611 | -105.5451
12.0648 | -105.5573
12.0686 | -105.5712
12.0724 | -105.5851
12.0761 | -105.5949
12.0799 | -105.6057
12.0836 | -105.6107
12.0874 | -105.6095
12.0911 | -105.6093
12.0949 | -105.6071
12.0987 | -105.5973
12.1024 | -105.5884
12.1062 | -105.5783
12.1099 | -105.5634
12.1137 | -105.5478
12.1175 | -105.5303
12.1212 | -105.5095
12.125 | -105.4878
12.1287 | -105.4691
12.1325 | -105.453
12.1362 | -105.433
12.14 | -105.4112
12.1438 | -105.3942
12.1475 | -105.3821
12.1513 | -105.3719
12.155 | -105.3682
12.1588 | -105.3724
12.1626 | -105.3735
12.1663 | -105.3759
12.1701 | -105.3849
12.1738 | -105.3954
12.1776 | -105.4102
12.1813 | -105.427
12.1851 | -105.4468
12.1889 | -105.4688
12.1926 | -105.4924
12.1964 | -105.5155
12.2001 | -105.5407
12.2039 | -105.5643
12.2077 | -105.5859
12.2114 | -105.6102
12.2152 | -105.6336
12.2189 | -105.6547
12.2227 | -105.676
12.2265 | -105.6936
12.2302 | -105.7087
12.234 | -105.7263
12.2377 | -105.7415
12.2415 | -105.751
12.2452 | -105.7555
12.249 | -105.7591
12.2528 | -105.7652
12.2565 | -105.768
12.2603 | -105.7732
12.264 | -105.7788
12.2678 | -105.7751
12.2716 | -105.7683
12.2753 | -105.7698
12.2791 | -105.7712
12.2828 | -105.7753
12.2866 | -105.7832
12.2903 | -105.7877
12.2941 | -105.7935
12.2979 | -105.7993
12.3016 | -105.8021
12.3054 | -105.8055
12.3091 | -105.8087
12.3129 | -105.8143
12.3167 | -105.8231
12.3204 | -105.8311
12.3242 | -105.8422
12.3279 | -105.8491
12.3317 | -105.8567
12.3354 | -105.8617
12.3392 | -105.8696
12.343 | -105.8819
12.3467 | -105.8884
12.3505 | -105.8935
12.3542 | -105.8975
12.358 | -105.9007
12.3618 | -105.9051
12.3655 | -105.9118
12.3693 | -105.9176
12.373 | -105.9217
12.3768 | -105.9239
12.3805 | -105.9262
12.3843 | -105.9305
12.3881 | -105.9303
12.3918 | -105.9325
12.3956 | -105.9358
12.3993 | -105.9382
12.4031 | -105.94
12.4069 | -105.9416
12.4106 | -105.9435
12.4144 | -105.9446
12.4181 | -105.9486
12.4219 | -105.9545
12.4257 | -105.9591
12.4294 | -105.963
12.4332 | -105.9661
12.4369 | -105.9686
12.4407 | -105.9746
12.4444 | -105.9875
12.4482 | -106.0002
12.452 | -106.003
12.4557 | -106.0052
12.4595 | -106.008
12.4632 | -106.0139
12.467 | -106.0197
12.4708 | -106.0228
12.4745 | -106.0269
12.4783 | -106.0316
12.482 | -106.0389
12.4858 | -106.0442
12.4895 | -106.0506
12.4933 | -106.0538
12.4971 | -106.055
12.5008 | -106.0569
12.5046 | -106.0572
12.5083 | -106.0618
12.5121 | -106.0637
12.5159 | -106.0694
12.5196 | -106.0753
12.5234 | -106.0791
12.5271 | -106.0827
12.5309 | -106.0863
12.5346 | -106.0897
12.5384 | -106.0943
12.5422 | -106.0931
12.5459 | -106.0902
12.5497 | -106.088
12.5534 | -106.086
12.5572 | -106.0776
12.561 | -106.0678
12.5647 | -106.0597
12.5685 | -106.051
12.5722 | -106.0408
12.576 | -106.0306
12.5798 | -106.0197
12.5835 | -106.0123
12.5873 | -106.0034
12.591 | -105.9899
12.5948 | -105.974
12.5985 | -105.9592
12.6023 | -105.9469
12.6061 | -105.9355
12.6098 | -105.9247
12.6136 | -105.9141
12.6173 | -105.9023
12.6211 | -105.8909
12.6249 | -105.8793
12.6286 | -105.8737
12.6324 | -105.8697
12.6361 | -105.8606
12.6399 | -105.8513
12.6436 | -105.8433
12.6474 | -105.8385
12.6512 | -105.8356
12.6549 | -105.8312
12.6587 | -105.8277
12.6624 | -105.8208
12.6662 | -105.8098
12.67 | -105.8002
12.6737 | -105.7893
12.6775 | -105.7798
12.6812 | -105.7726
12.685 | -105.7632
12.6887 | -105.7536
12.6925 | -105.7433
12.6963 | -105.7351
12.7 | -105.7267
12.7038 | -105.7209
12.7075 | -105.717
12.7113 | -105.7165
12.7151 | -105.7173
12.7188 | -105.7193
12.7226 | -105.7213
12.7263 | -105.7277
12.7301 | -105.7383
12.7339 | -105.7509
12.7376 | -105.7644
12.7414 | -105.78
12.7451 | -105.7954
12.7489 | -105.8064
12.7526 | -105.8209
12.7564 | -105.8364
12.7602 | -105.8517
12.7639 | -105.8655
12.7677 | -105.8775
12.7714 | -105.8883
12.7752 | -105.9
12.779 | -105.9088
12.7827 | -105.9136
12.7865 | -105.9182
12.7902 | -105.9172
12.794 | -105.9096
12.7977 | -105.9014
12.8015 | -105.8918
12.8053 | -105.8755
12.809 | -105.8571
12.8128 | -105.8308
12.8165 | -105.8088
12.8203 | -105.7895
12.8241 | -105.7661
12.8278 | -105.7417
12.8316 | -105.7184
12.8353 | -105.6946
12.8391 | -105.6739
12.8428 | -105.6569
12.8466 | -105.6445
12.8504 | -105.6377
12.8541 | -105.6289
12.8579 | -105.6228
12.8616 | -105.6212
12.8654 | -105.6209
12.8692 | -105.6266
12.8729 | -105.6373
12.8767 | -105.6523
12.8804 | -105.6683
12.8842 | -105.686
12.8879 | -105.7037
12.8917 | -105.7247
12.8955 | -105.748
12.8992 | -105.7743
12.903 | -105.8
12.9067 | -105.825
12.9105 | -105.8471
12.9143 | -105.8658
12.918 | -105.8863
12.9218 | -105.9073
12.9255 | -105.9287
12.9293 | -105.9455
12.9331 | -105.9612
12.9368 | -105.9716
12.9406 | -105.9785
12.9443 | -105.9838
12.9481 | -105.9853
12.9518 | -105.9864
12.9556 | -105.9896
12.9594 | -105.9914
12.9631 | -105.9938
12.9669 | -105.9957
12.9706 | -105.9941
12.9744 | -105.9972
12.9782 | -105.9997
12.9819 | -106.0027
12.9857 | -106.0078
12.9894 | -106.0169
12.9932 | -106.0225
12.9969 | -106.0323
13.0007 | -106.0418
13.0045 | -106.0506
13.0082 | -106.0616
13.012 | -106.0732
13.0157 | -106.0829
13.0195 | -106.0916
13.0233 | -106.0976
13.027 | -106.1
13.0308 | -106.1021
13.0345 | -106.102
13.0383 | -106.1056
13.042 | -106.1123
13.0458 | -106.1175
13.0496 | -106.1188
13.0533 | -106.1196
13.0571 | -106.119
13.0608 | -106.1205
13.0646 | -106.122
13.0684 | -106.1223
13.0721 | -106.1227
13.0759 | -106.1274
13.0796 | -106.1308
13.0834 | -106.133
13.0872 | -106.1341
13.0909 | -106.1366
13.0947 | -106.1425
13.0984 | -106.1491
13.1022 | -106.1516
13.1059 | -106.152
13.1097 | -106.1529
13.1135 | -106.1569
13.1172 | -106.1597
13.121 | -106.1591
13.1247 | -106.1586
13.1285 | -106.1572
13.1323 | -106.1551
13.136 | -106.1566
13.1398 | -106.1568
13.1435 | -106.1536
13.1473 | -106.1517
13.151 | -106.1497
13.1548 | -106.1472
13.1586 | -106.1466
13.1623 | -106.1473
13.1661 | -106.1531
13.1698 | -106.1594
13.1736 | -106.1658
13.1774 | -106.1713
13.1811 | -106.1796
13.1849 | -106.1862
13.1886 | -106.1928
13.1924 | -106.2005
13.1961 | -106.2108
13.1999 | -106.2229
13.2037 | -106.2357
13.2074 | -106.2494
13.2112 | -106.265
13.2149 | -106.2853
13.2187 | -106.3076
13.2225 | -106.3307
13.2262 | -106.3547
13.23 | -106.3726
13.2337 | -106.389
13.2375 | -106.4065
13.2413 | -106.425
13.245 | -106.4404
13.2488 | -106.4588
13.2525 | -106.474
13.2563 | -106.4854
13.26 | -106.4957
13.2638 | -106.5033
13.2676 | -106.5087
13.2713 | -106.5111
13.2751 | -106.5134
13.2788 | -106.5178
13.2826 | -106.5222
13.2864 | -106.5223
13.2901 | -106.5225
13.2939 | -106.5224
13.2976 | -106.5204
13.3014 | -106.5178
13.3051 | -106.5096
13.3089 | -106.4987
13.3127 | -106.4913
13.3164 | -106.4867
13.3202 | -106.4837
13.3239 | -106.4818
13.3277 | -106.4772
13.3315 | -106.467
13.3352 | -106.4588
13.339 | -106.4492
13.3427 | -106.4411
13.3465 | -106.4366
13.3502 | -106.4301
13.354 | -106.4189
13.3578 | -106.4096
13.3615 | -106.4034
13.3653 | -106.398
13.369 | -106.3949
13.3728 | -106.3951
13.3766 | -106.3965
13.3803 | -106.395
13.3841 | -106.3909
13.3878 | -106.3815
13.3916 | -106.3737
13.3953 | -106.3699
13.3991 | -106.3658
13.4029 | -106.3614
13.4066 | -106.3585
13.4104 | -106.3562
13.4141 | -106.3533
13.4179 | -106.3494
13.4217 | -106.3458
13.4254 | -106.3435
13.4292 | -106.3419
13.4329 | -106.3417
13.4367 | -106.3422
13.4405 | -106.3396
13.4442 | -106.3395
13.448 | -106.3425
13.4517 | -106.3473
13.4555 | -106.3528
13.4592 | -106.3557
13.463 | -106.3596
13.4668 | -106.3623
13.4705 | -106.3641
13.4743 | -106.3649
13.478 | -106.3672
13.4818 | -106.3692
13.4856 | -106.3736
13.4893 | -106.3784
13.4931 | -106.3811
13.4968 | -106.3883
13.5006 | -106.3976
13.5043 | -106.4037
13.5081 | -106.4065
13.5119 | -106.4063
13.5156 | -106.405
13.5194 | -106.4059
13.5231 | -106.4103
13.5269 | -106.4128
13.5307 | -106.4136
13.5344 | -106.4105
13.5382 | -106.404
13.5419 | -106.4004
13.5457 | -106.3972
13.5494 | -106.3862
13.5532 | -106.3732
13.557 | -106.3624
13.5607 | -106.3483
13.5645 | -106.334
13.5682 | -106.3187
13.572 | -106.3015
13.5758 | -106.2848
13.5795 | -106.2729
13.5833 | -106.264
13.587 | -106.2479
13.5908 | -106.2286
13.5945 | -106.2121
13.5983 | -106.1942
13.6021 | -106.1762
13.6058 | -106.1576
13.6096 | -106.1404
13.6133 | -106.1244
13.6171 | -106.1095
13.6209 | -106.0968
13.6246 | -106.0836
13.6284 | -106.0699
13.6321 | -106.0553
13.6359 | -106.0404
13.6397 | -106.0282
13.6434 | -106.0178
13.6472 | -106.007
13.6509 | -105.9968
13.6547 | -105.9885
13.6584 | -105.9797
13.6622 | -105.9715
13.666 | -105.9636
13.6697 | -105.9559
13.6735 | -105.9484
13.6772 | -105.9414
13.681 | -105.9332
13.6848 | -105.9249
13.6885 | -105.9176
13.6923 | -105.9092
13.696 | -105.9023
13.6998 | -105.8956
13.7035 | -105.8883
13.7073 | -105.8844
13.7111 | -105.8813
13.7148 | -105.8776
13.7186 | -105.874
13.7223 | -105.8703
13.7261 | -105.8704
13.7299 | -105.8738
13.7336 | -105.8769
13.7374 | -105.8787
13.7411 | -105.8827
13.7449 | -105.8897
13.7486 | -105.8964
13.7524 | -105.9032
13.7562 | -105.9113
13.7599 | -105.924
13.7637 | -105.9373
13.7674 | -105.9508
13.7712 | -105.9676
13.775 | -105.9841
13.7787 | -105.9972
13.7825 | -106.0114
13.7862 | -106.0296
13.79 | -106.0462
13.7938 | -106.0586
13.7975 | -106.0749
13.8013 | -106.0907
13.805 | -106.105
13.8088 | -106.1209
13.8125 | -106.1375
13.8163 | -106.1536
13.8201 | -106.1629
13.8238 | -106.1738
13.8276 | -106.1858
13.8313 | -106.1962
13.8351 | -106.2039
13.8389 | -106.2117
13.8426 | -106.216
13.8464 | -106.2185
13.8501 | -106.221
13.8539 | -106.2238
13.8576 | -106.2287
13.8614 | -106.232
13.8652 | -106.2337
13.8689 | -106.237
13.8727 | -106.24
13.8764 | -106.2418
13.8802 | -106.2445
13.884 | -106.2457
13.8877 | -106.2449
13.8915 | -106.2459
13.8952 | -106.2453
13.899 | -106.2456
13.9027 | -106.2466
13.9065 | -106.2474
13.9103 | -106.2462
13.914 | -106.2448
13.9178 | -106.2449
13.9215 | -106.2447
13.9253 | -106.2415
13.9291 | -106.2356
13.9328 | -106.2304
13.9366 | -106.2235
13.9403 | -106.2157
13.9441 | -106.2053
13.9479 | -106.193
13.9516 | -106.1793
13.9554 | -106.166
13.9591 | -106.1505
13.9629 | -106.1318
13.9666 | -106.1109
13.9704 | -106.0889
13.9742 | -106.0669
13.9779 | -106.0428
13.9817 | -106.0164
13.9854 | -105.9903
13.9892 | -105.9629
13.993 | -105.937
13.9967 | -105.909
14.0005 | -105.8807
14.0042 | -105.8523
14.008 | -105.8199
14.0117 | -105.789
14.0155 | -105.7618
14.0193 | -105.7389
14.023 | -105.714
14.0268 | -105.6887
14.0305 | -105.6676
14.0343 | -105.6485
14.0381 | -105.6282
14.0418 | -105.6147
14.0456 | -105.604
14.0493 | -105.5931
14.0531 | -105.5835
14.0568 | -105.5774
14.0606 | -105.5756
14.0644 | -105.5784
14.0681 | -105.5827
14.0719 | -105.5905
14.0756 | -105.6007
14.0794 | -105.6126
14.0832 | -105.6285
14.0869 | -105.6465
14.0907 | -105.6666
14.0944 | -105.6906
14.0982 | -105.7123
14.102 | -105.7335
14.1057 | -105.756
14.1095 | -105.7784
14.1132 | -105.8032
14.117 | -105.8277
14.1207 | -105.8501
14.1245 | -105.8708
14.1283 | -105.8914
14.132 | -105.9157
14.1358 | -105.9401
14.1395 | -105.9654
14.1433 | -105.9882
14.1471 | -106.0115
14.1508 | -106.0333
14.1546 | -106.0558
14.1583 | -106.0763
14.1621 | -106.0958
14.1658 | -106.1166
14.1696 | -106.1392
14.1734 | -106.1611
14.1771 | -106.1827
14.1809 | -106.2073
14.1846 | -106.2303
14.1884 | -106.2498
14.1922 | -106.2737
14.1959 | -106.2978
14.1997 | -106.3176
14.2034 | -106.3353
14.2072 | -106.3555
14.2109 | -106.3761
14.2147 | -106.3912
14.2185 | -106.4031
14.2222 | -106.413
14.226 | -106.4218
14.2297 | -106.4278
14.2335 | -106.4313
14.2373 | -106.4325
14.241 | -106.4329
14.2448 | -106.4276
14.2485 | -106.4202
14.2523 | -106.4119
14.256 | -106.4036
14.2598 | -106.39
14.2636 | -106.3749
14.2673 | -106.3583
14.2711 | -106.3422
14.2748 | -106.3253
14.2786 | -106.3098
14.2824 | -106.2946
14.2861 | -106.28
14.2899 | -106.2673
14.2936 | -106.2521
14.2974 | -106.243
14.3012 | -106.2367
14.3049 | -106.2314
14.3087 | -106.2291
14.3124 | -106.2282
14.3162 | -106.2323
14.3199 | -106.2396
14.3237 | -106.2505
14.3275 | -106.2649
14.3312 | -106.2796
14.335 | -106.297
14.3387 | -106.3137
14.3425 | -106.3325
14.3463 | -106.3536
14.35 | -106.3772
14.3538 | -106.402
14.3575 | -106.4268
14.3613 | -106.4506
14.365 | -106.4772
14.3688 | -106.4995
14.3726 | -106.523
14.3763 | -106.5454
14.3801 | -106.5637
14.3838 | -106.5793
14.3876 | -106.5929
14.3914 | -106.6048
14.3951 | -106.6139
14.3989 | -106.6173
14.4026 | -106.6187
14.4064 | -106.6185
14.4101 | -106.6191
14.4139 | -106.6119
14.4177 | -106.6026
14.4214 | -106.5942
14.4252 | -106.5829
14.4289 | -106.5689
14.4327 | -106.5524
14.4365 | -106.5358
14.4402 | -106.5188
14.444 | -106.4997
14.4477 | -106.4773
14.4515 | -106.4549
14.4552 | -106.4368
14.459 | -106.4174
14.4628 | -106.3996
14.4665 | -106.3851
14.4703 | -106.3708
14.474 | -106.3549
14.4778 | -106.3402
14.4816 | -106.3259
14.4853 | -106.3131
14.4891 | -106.2993
14.4928 | -106.2863
14.4966 | -106.2746
14.5004 | -106.2666
14.5041 | -106.2571
14.5079 | -106.2488
14.5116 | -106.2422
14.5154 | -106.236
14.5191 | -106.2321
14.5229 | -106.2243
14.5267 | -106.2193
14.5304 | -106.2147
14.5342 | -106.2114
14.5379 | -106.2117
14.5417 | -106.2128
14.5455 | -106.2165
14.5492 | -106.222
14.553 | -106.2277
14.5567 | -106.2359
14.5605 | -106.2449
14.5642 | -106.2525
14.568 | -106.2641
14.5718 | -106.2755
14.5755 | -106.2896
14.5793 | -106.2995
14.583 | -106.308
14.5868 | -106.3193
14.5906 | -106.332
14.5943 | -106.341
14.5981 | -106.3474
14.6018 | -106.3562
14.6056 | -106.3652
14.6093 | -106.3698
14.6131 | -106.3742
14.6169 | -106.3763
14.6206 | -106.3813
14.6244 | -106.3851
14.6281 | -106.3872
14.6319 | -106.3856
14.6357 | -106.3843
14.6394 | -106.3783
14.6432 | -106.3775
14.6469 | -106.3747
14.6507 | -106.3688
14.6545 | -106.3614
14.6582 | -106.3549
14.662 | -106.3486
14.6657 | -106.339
14.6695 | -106.3293
14.6732 | -106.3177
14.677 | -106.3064
14.6808 | -106.2932
14.6845 | -106.2811
14.6883 | -106.2701
14.692 | -106.2579
14.6958 | -106.2444
14.6996 | -106.2338
14.7033 | -106.2215
14.7071 | -106.2082
14.7108 | -106.1931
14.7146 | -106.1803
14.7183 | -106.1681
14.7221 | -106.155
14.7259 | -106.1454
14.7296 | -106.1376
14.7334 | -106.1342
14.7371 | -106.1316
14.7409 | -106.1303
14.7447 | -106.1283
14.7484 | -106.1284
14.7522 | -106.1303
14.7559 | -106.1325
14.7597 | -106.1369
14.7634 | -106.1461
14.7672 | -106.1569
14.771 | -106.1698
14.7747 | -106.1858
14.7785 | -106.2029
14.7822 | -106.2213
14.786 | -106.2424
14.7898 | -106.2674
14.7935 | -106.2926
14.7973 | -106.3152
14.801 | -106.3353
14.8048 | -106.3592
14.8086 | -106.3852
14.8123 | -106.4082
14.8161 | -106.4271
14.8198 | -106.4481
14.8236 | -106.47
14.8273 | -106.4889
14.8311 | -106.5062
14.8349 | -106.5204
14.8386 | -106.5329
14.8424 | -106.5473
14.8461 | -106.5601
14.8499 | -106.5692
14.8537 | -106.5764
14.8574 | -106.5836
14.8612 | -106.5898
14.8649 | -106.5962
14.8687 | -106.6043
14.8724 | -106.6065
14.8762 | -106.6073
14.88 | -106.6064
14.8837 | -106.6021
14.8875 | -106.5961
14.8912 | -106.5905
14.895 | -106.5868
14.8988 | -106.5815
14.9025 | -106.576
14.9063 | -106.5697
14.91 | -106.5662
14.9138 | -106.5619
14.9175 | -106.5573
14.9213 | -106.553
14.9251 | -106.5501
14.9288 | -106.5456
14.9326 | -106.5388
14.9363 | -106.5336
14.9401 | -106.5268
14.9439 | -106.5198
14.9476 | -106.5134
14.9514 | -106.5071
14.9551 | -106.5015
14.9589 | -106.4972
14.9627 | -106.4914
14.9664 | -106.4832
14.9702 | -106.4743
14.9739 | -106.4641
14.9777 | -106.4535
14.9814 | -106.444
14.9852 | -106.4346
14.989 | -106.4249
14.9927 | -106.4141
14.9965 | -106.4053
15.0002 | -106.3979
15.004 | -106.3884
15.0078 | -106.3806
15.0115 | -106.3744
15.0153 | -106.3714
15.019 | -106.3678
15.0228 | -106.3653
15.0265 | -106.3638
15.0303 | -106.3646
15.0341 | -106.3647
15.0378 | -106.3656
15.0416 | -106.3684
15.0453 | -106.3716
15.0491 | -106.3774
15.0529 | -106.3836
15.0566 | -106.3895
15.0604 | -106.3955
15.0641 | -106.4014
15.0679 | -106.4085
15.0716 | -106.4156
15.0754 | -106.4219
15.0792 | -106.4265
15.0829 | -106.4314
15.0867 | -106.4413
15.0904 | -106.4501
15.0942 | -106.458
15.098 | -106.4689
15.1017 | -106.4769
15.1055 | -106.4832
15.1092 | -106.4885
15.113 | -106.4946
15.1167 | -106.497
15.1205 | -106.4992
15.1243 | -106.5009
15.128 | -106.5016
15.1318 | -106.5012
15.1355 | -106.5014
15.1393 | -106.5011
15.1431 | -106.4998
15.1468 | -106.5005
15.1506 | -106.4992
15.1543 | -106.4992
15.1581 | -106.4992
15.1619 | -106.495
15.1656 | -106.4902
15.1694 | -106.4885
15.1731 | -106.4874
15.1769 | -106.4869
15.1806 | -106.4887
15.1844 | -106.4884
15.1882 | -106.4897
15.1919 | -106.4903
15.1957 | -106.489
15.1994 | -106.488
15.2032 | -106.487
15.207 | -106.4863
15.2107 | -106.4855
15.2145 | -106.4876
15.2182 | -106.4903
15.222 | -106.4952
15.2257 | -106.5006
15.2295 | -106.5067
15.2333 | -106.513
15.237 | -106.5178
15.2408 | -106.5249
15.2445 | -106.5332
15.2483 | -106.5426
15.2521 | -106.5547
15.2558 | -106.5654
15.2596 | -106.5761
15.2633 | -106.587
15.2671 | -106.5978
15.2708 | -106.6096
15.2746 | -106.6191
15.2784 | -106.629
15.2821 | -106.6386
15.2859 | -106.6453
15.2896 | -106.6501
15.2934 | -106.6586
15.2972 | -106.6656
15.3009 | -106.6658
15.3047 | -106.6641
15.3084 | -106.662
15.3122 | -106.6597
15.316 | -106.6554
15.3197 | -106.6509
15.3235 | -106.6423
15.3272 | -106.636
15.331 | -106.6277
15.3347 | -106.6212
15.3385 | -106.6154
15.3423 | -106.6052
15.346 | -106.5955
15.3498 | -106.5871
15.3535 | -106.5793
15.3573 | -106.5698
15.3611 | -106.5609
15.3648 | -106.5556
15.3686 | -106.5514
15.3723 | -106.5456
15.3761 | -106.5415
15.3798 | -106.5387
15.3836 | -106.5355
15.3874 | -106.5342
15.3911 | -106.5327
15.3949 | -106.5304
15.3986 | -106.5276
15.4024 | -106.5252
15.4062 | -106.5264
15.4099 | -106.5276
15.4137 | -106.529
15.4174 | -106.5306
15.4212 | -106.5318
15.4249 | -106.5297
15.4287 | -106.5308
15.4325 | -106.5301
15.4362 | -106.5245
15.44 | -106.5164
15.4437 | -106.5087
15.4475 | -106.4995
15.4513 | -106.4924
15.455 | -106.4845
15.4588 | -106.4735
15.4625 | -106.457
15.4663 | -106.4444
15.47 | -106.4303
15.4738 | -106.4165
15.4776 | -106.4023
15.4813 | -106.3861
15.4851 | -106.3709
15.4888 | -106.3557
15.4926 | -106.3398
15.4964 | -106.3252
15.5001 | -106.3114
15.5039 | -106.2977
15.5076 | -106.2842
15.5114 | -106.273
15.5152 | -106.2601
15.5189 | -106.2494
15.5227 | -106.2421
15.5264 | -106.234
15.5302 | -106.2285
15.5339 | -106.2273
15.5377 | -106.2279
15.5415 | -106.2288
15.5452 | -106.2329
15.549 | -106.2393
15.5527 | -106.25
15.5565 | -106.261
15.5603 | -106.2724
15.564 | -106.2886
15.5678 | -106.3076
15.5715 | -106.3273
15.5753 | -106.3445
15.579 | -106.3649
15.5828 | -106.3879
15.5866 | -106.4046
15.5903 | -106.4234
15.5941 | -106.4453
15.5978 | -106.4672
15.6016 | -106.4884
15.6054 | -106.5081
15.6091 | -106.5248
15.6129 | -106.5413
15.6166 | -106.5565
15.6204 | -106.5702
15.6241 | -106.5827
15.6279 | -106.5917
15.6317 | -106.5978
15.6354 | -106.6012
15.6392 | -106.6049
15.6429 | -106.6047
15.6467 | -106.6064
15.6505 | -106.6019
15.6542 | -106.5975
15.658 | -106.5937
15.6617 | -106.5889
15.6655 | -106.5805
15.6693 | -106.5731
15.673 | -106.5627
15.6768 | -106.5509
15.6805 | -106.5413
15.6843 | -106.5323
15.688 | -106.5197
15.6918 | -106.5076
15.6956 | -106.4957
15.6993 | -106.4861
15.7031 | -106.4797
15.7068 | -106.4742
15.7106 | -106.4701
15.7144 | -106.4683
15.7181 | -106.47
15.7219 | -106.487
15.7256 | -106.4923
15.7294 | -106.4973
15.7331 | -106.5066
15.7369 | -106.5153
15.7407 | -106.5246
15.7444 | -106.5345
15.7482 | -106.5458
15.7519 | -106.5615
15.7557 | -106.5772
15.7595 | -106.59
15.7632 | -106.6046
15.767 | -106.6219
15.7707 | -106.6372
15.7745 | -106.6481
15.7782 | -106.6566
15.782 | -106.6663
15.7858 | -106.6735
15.7895 | -106.6798
15.7933 | -106.684
15.797 | -106.6882
15.8008 | -106.6911
15.8046 | -106.689
15.8083 | -106.6849
15.8121 | -106.6854
15.8158 | -106.6806
15.8196 | -106.6715
15.8233 | -106.6613
15.8271 | -106.652
15.8309 | -106.6408
15.8346 | -106.6254
15.8384 | -106.6103
15.8421 | -106.5966
15.8459 | -106.5821
15.8497 | -106.5642
15.8534 | -106.547
15.8572 | -106.531
15.8609 | -106.5166
15.8647 | -106.5018
15.8685 | -106.4894
15.8722 | -106.4766
15.876 | -106.4662
15.8797 | -106.4561
15.8835 | -106.4467
15.8872 | -106.4386
15.891 | -106.4318
15.8948 | -106.4263
15.8985 | -106.4236
15.9023 | -106.4224
15.906 | -106.4237
15.9098 | -106.4259
15.9136 | -106.4271
15.9173 | -106.4292
15.9211 | -106.4318
15.9248 | -106.4359
15.9286 | -106.4414
15.9323 | -106.4459
15.9361 | -106.4496
15.9399 | -106.453
15.9436 | -106.4552
15.9474 | -106.4568
15.9511 | -106.457
15.9549 | -106.4587
15.9587 | -106.4625
15.9624 | -106.462
15.9662 | -106.4607
15.9699 | -106.4618
15.9737 | -106.4626
15.9774 | -106.4606
15.9812 | -106.4554
15.985 | -106.4483
15.9887 | -106.4416
15.9925 | -106.4342
15.9962 | -106.4264
16 | -106.4189
16.0038 | -106.4104
16.0075 | -106.4023
16.0113 | -106.393
16.015 | -106.3865
16.0188 | -106.3771
16.0226 | -106.3644
16.0263 | -106.3521
16.0301 | -106.3431
16.0338 | -106.3361
16.0376 | -106.3274
16.0413 | -106.3176
16.0451 | -106.3103
16.0489 | -106.3037
16.0526 | -106.2966
16.0564 | -106.2908
16.0601 | -106.2868
16.0639 | -106.2845
16.0677 | -106.2808
16.0714 | -106.2759
16.0752 | -106.2719
16.0789 | -106.2697
16.0827 | -106.2684
16.0864 | -106.2666
16.0902 | -106.2669
16.094 | -106.2667
16.0977 | -106.2657
16.1015 | -106.2654
16.1052 | -106.2673
16.109 | -106.272
16.1128 | -106.2761
16.1165 | -106.2815
16.1203 | -106.289
16.124 | -106.2957
16.1278 | -106.3007
16.1315 | -106.3061
16.1353 | -106.3108
16.1391 | -106.3162
16.1428 | -106.324
16.1466 | -106.3328
16.1503 | -106.3414
16.1541 | -106.3481
16.1579 | -106.3522
16.1616 | -106.3598
16.1654 | -106.3675
16.1691 | -106.3728
16.1729 | -106.3756
16.1766 | -106.3793
16.1804 | -106.3839
16.1842 | -106.3854
16.1879 | -106.3843
16.1917 | -106.3832
16.1954 | -106.3794
16.1992 | -106.3746
16.203 | -106.3681
16.2067 | -106.3608
16.2105 | -106.3522
16.2142 | -106.342
16.218 | -106.33
16.2218 | -106.3173
16.2255 | -106.3041
16.2293 | -106.2897
16.233 | -106.2737
16.2368 | -106.2577
16.2405 | -106.2415
16.2443 | -106.2229
16.2481 | -106.2041
16.2518 | -106.188
16.2556 | -106.1695
16.2593 | -106.1483
16.2631 | -106.1289
16.2669 | -106.1121
16.2706 | -106.0943
16.2744 | -106.0761
16.2781 | -106.0596
16.2819 | -106.0431
16.2856 | -106.0274
16.2894 | -106.0142
16.2932 | -106.0015
16.2969 | -105.9889
16.3007 | -105.9796
16.3044 | -105.9708
16.3082 | -105.9614
16.312 | -105.9549
16.3157 | -105.9487
16.3195 | -105.9447
16.3232 | -105.9423
16.327 | -105.9408
16.3307 | -105.9407
16.3345 | -105.9435
16.3383 | -105.9474
16.342 | -105.9507
16.3458 | -105.9564
16.3495 | -105.9637
16.3533 | -105.9733
16.3571 | -105.9848
16.3608 | -105.997
16.3646 | -106.0105
16.3683 | -106.0272
16.3721 | -106.0452
16.3759 | -106.0616
16.3796 | -106.0767
16.3834 | -106.0929
16.3871 | -106.1112
16.3909 | -106.1304
16.3946 | -106.1459
16.3984 | -106.162
16.4022 | -106.1785
16.4059 | -106.195
16.4097 | -106.2097
16.4134 | -106.222
16.4172 | -106.2343
16.421 | -106.2464
16.4247 | -106.257
16.4285 | -106.2676
16.4322 | -106.2747
16.436 | -106.2786
16.4397 | -106.2829
16.4435 | -106.2843
16.4473 | -106.281
16.451 | -106.2805
16.4548 | -106.2774
16.4585 | -106.2699
16.4623 | -106.2655
16.4661 | -106.2624
16.4698 | -106.2569
16.4736 | -106.2493
16.4773 | -106.2437
16.4811 | -106.2374
16.4848 | -106.2309
16.4886 | -106.2242
16.4924 | -106.2198
16.4961 | -106.2136
16.4999 | -106.2062
16.5036 | -106.2015
16.5074 | -106.198
16.5112 | -106.196
16.5149 | -106.1952
16.5187 | -106.1983
16.5224 | -106.201
16.5262 | -106.2047
16.53 | -106.2089
16.5337 | -106.2148
16.5375 | -106.2207
16.5412 | -106.2309
16.545 | -106.243
16.5487 | -106.2544
16.5525 | -106.2676
16.5563 | -106.2818
16.56 | -106.2961
16.5638 | -106.3129
16.5675 | -106.3324
16.5713 | -106.3496
16.5751 | -106.3669
16.5788 | -106.3855
16.5826 | -106.4059
16.5863 | -106.427
16.5901 | -106.445
16.5938 | -106.4612
16.5976 | -106.4812
16.6014 | -106.5007
16.6051 | -106.5178
16.6089 | -106.5327
16.6126 | -106.5489
16.6164 | -106.5644
16.6202 | -106.5783
16.6239 | -106.5903
16.6277 | -106.6036
16.6314 | -106.6159
16.6352 | -106.6267
16.6389 | -106.6374
16.6427 | -106.6462
16.6465 | -106.6537
16.6502 | -106.66
16.654 | -106.663
16.6577 | -106.6683
16.6615 | -106.6729
16.6653 | -106.675
16.669 | -106.6742
16.6728 | -106.675
16.6765 | -106.6757
16.6803 | -106.677
16.684 | -106.6808
16.6878 | -106.6789
16.6916 | -106.6784
16.6953 | -106.6805
16.6991 | -106.6799
16.7028 | -106.6777
16.7066 | -106.6757
16.7104 | -106.673
16.7141 | -106.6689
16.7179 | -106.6662
16.7216 | -106.6616
16.7254 | -106.6598
16.7292 | -106.6564
16.7329 | -106.6515
16.7367 | -106.6465
16.7404 | -106.6414
16.7442 | -106.635
16.7479 | -106.6279
16.7517 | -106.6209
16.7555 | -106.6117
16.7592 | -106.6019
16.763 | -106.5913
16.7667 | -106.5815
16.7705 | -106.571
16.7743 | -106.5609
16.778 | -106.5518
16.7818 | -106.5441
16.7855 | -106.5357
16.7893 | -106.528
16.793 | -106.5208
16.7968 | -106.513
16.8006 | -106.5013
16.8043 | -106.4895
16.8081 | -106.4854
16.8118 | -106.4794
16.8156 | -106.472
16.8194 | -106.4662
16.8231 | -106.4611
16.8269 | -106.4556
16.8306 | -106.4517
16.8344 | -106.4478
16.8381 | -106.4432
16.8419 | -106.4406
16.8457 | -106.4386
16.8494 | -106.4355
16.8532 | -106.4331
16.8569 | -106.4331
16.8607 | -106.4311
16.8645 | -106.4306
16.8682 | -106.4296
16.872 | -106.4316
16.8757 | -106.434
16.8795 | -106.436
16.8833 | -106.439
16.887 | -106.4415
16.8908 | -106.4469
16.8945 | -106.4547
16.8983 | -106.4612
16.902 | -106.4671
16.9058 | -106.4726
16.9096 | -106.4795
16.9133 | -106.4856
16.9171 | -106.4907
16.9208 | -106.4958
16.9246 | -106.4986
16.9284 | -106.5028
16.9321 | -106.5051
16.9359 | -106.5047
16.9396 | -106.5022
16.9434 | -106.5023
16.9471 | -106.5028
16.9509 | -106.504
16.9547 | -106.5026
16.9584 | -106.4987
16.9622 | -106.4959
16.9659 | -106.4918
16.9697 | -106.4871
16.9735 | -106.4822
16.9772 | -106.4776
16.981 | -106.4742
16.9847 | -106.4707
16.9885 | -106.4676
16.9922 | -106.4643
16.996 | -106.462
16.9998 | -106.4617
17.0035 | -106.4612
17.0073 | -106.4622
17.011 | -106.4656
17.0148 | -106.4667
17.0186 | -106.4663
17.0223 | -106.4646
17.0261 | -106.4645
17.0298 | -106.4659
17.0336 | -106.4682
17.0373 | -106.4743
17.0411 | -106.4863
17.0449 | -106.4958
17.0486 | -106.5005
17.0524 | -106.5048
17.0561 | -106.5094
17.0599 | -106.5134
17.0637 | -106.5186
17.0674 | -106.5234
17.0712 | -106.5295
17.0749 | -106.5355
17.0787 | -106.5415
17.0825 | -106.5461
17.0862 | -106.5499
17.09 | -106.5537
17.0937 | -106.5583
17.0975 | -106.5618
17.1012 | -106.5645
17.105 | -106.5682
17.1088 | -106.5701
17.1125 | -106.5738
17.1163 | -106.5777
17.12 | -106.5811
17.1238 | -106.5859
17.1276 | -106.5913
17.1313 | -106.5974
17.1351 | -106.602
17.1388 | -106.6075
17.1426 | -106.6145
17.1463 | -106.6225
17.1501 | -106.6347
17.1539 | -106.6487
17.1576 | -106.663
17.1614 | -106.6764
17.1651 | -106.69
17.1689 | -106.7062
17.1727 | -106.7226
17.1764 | -106.7415
17.1802 | -106.7626
17.1839 | -106.7834
17.1877 | -106.8034
17.1914 | -106.8232
17.1952 | -106.8423
17.199 | -106.8627
17.2027 | -106.8847
17.2065 | -106.9056
17.2102 | -106.925
17.214 | -106.9418
17.2178 | -106.9558
17.2215 | -106.9711
17.2253 | -106.9826
17.229 | -106.9934
17.2328 | -107.0047
17.2366 | -107.0122
17.2403 | -107.0169
17.2441 | -107.0237
17.2478 | -107.0267
17.2516 | -107.0229
17.2553 | -107.0181
17.2591 | -107.0136
17.2629 | -107.0057
17.2666 | -106.9972
17.2704 | -106.9853
17.2741 | -106.9679
17.2779 | -106.9502
17.2817 | -106.9298
17.2854 | -106.9068
17.2892 | -106.8843
17.2929 | -106.8623
17.2967 | -106.8407
17.3004 | -106.8181
17.3042 | -106.7985
17.308 | -106.7744
17.3117 | -106.75
17.3155 | -106.7246
17.3192 | -106.6999
17.323 | -106.6797
17.3268 | -106.6621
17.3305 | -106.6421
17.3343 | -106.6238
17.338 | -106.6098
17.3418 | -106.598
17.3455 | -106.5886
17.3493 | -106.5918
17.3531 | -106.5869
17.3568 | -106.5833
17.3606 | -106.5792
17.3643 | -106.5768
17.3681 | -106.578
17.3719 | -106.5826
17.3756 | -106.5859
17.3794 | -106.5891
17.3831 | -106.5945
17.3869 | -106.6005
17.3907 | -106.607
17.3944 | -106.6133
17.3982 | -106.6194
17.4019 | -106.6271
17.4057 | -106.636
17.4094 | -106.6443
17.4132 | -106.652
17.417 | -106.6583
17.4207 | -106.6644
17.4245 | -106.6674
17.4282 | -106.6697
17.432 | -106.673
17.4358 | -106.6754
17.4395 | -106.6772
17.4433 | -106.6824
17.447 | -106.6866
17.4508 | -106.6884
17.4545 | -106.6911
17.4583 | -106.6934
17.4621 | -106.6992
17.4658 | -106.7057
17.4696 | -106.7089
17.4733 | -106.7086
17.4771 | -106.7104
17.4809 | -106.7107
17.4846 | -106.7155
17.4884 | -106.7221
17.4921 | -106.7258
17.4959 | -106.7319
17.4996 | -106.7428
17.5034 | -106.7502
17.5072 | -106.7575
17.5109 | -106.7672
17.5147 | -106.7781
17.5184 | -106.7884
17.5222 | -106.7994
17.526 | -106.8109
17.5297 | -106.8219
17.5335 | -106.8345
17.5372 | -106.8463
17.541 | -106.8579
17.5448 | -106.869
17.5485 | -106.8825
17.5523 | -106.8922
17.556 | -106.904
17.5598 | -106.9136
17.5635 | -106.9229
17.5673 | -106.934
17.5711 | -106.94
17.5748 | -106.9445
17.5786 | -106.9517
17.5823 | -106.9566
17.5861 | -106.9615
17.5899 | -106.9634
17.5936 | -106.9642
17.5974 | -106.9635
17.6011 | -106.9638
17.6049 | -106.9601
17.6086 | -106.9506
17.6124 | -106.9418
17.6162 | -106.9346
17.6199 | -106.9262
17.6237 | -106.9122
17.6274 | -106.9011
17.6312 | -106.8877
17.635 | -106.8733
17.6387 | -106.8585
17.6425 | -106.8408
17.6462 | -106.8237
17.65 | -106.8045
17.6537 | -106.7837
17.6575 | -106.7611
17.6613 | -106.7387
17.665 | -106.7185
17.6688 | -106.6989
17.6725 | -106.6742
17.6763 | -106.6499
17.6801 | -106.6269
17.6838 | -106.6035
17.6876 | -106.5816
17.6913 | -106.5583
17.6951 | -106.5329
17.6988 | -106.5093
17.7026 | -106.4887
17.7064 | -106.4689
17.7101 | -106.451
17.7139 | -106.4329
17.7176 | -106.4194
17.7214 | -106.4089
17.7252 | -106.4015
17.7289 | -106.3926
17.7327 | -106.3847
17.7364 | -106.3785
17.7402 | -106.3739
17.744 | -106.3706
17.7477 | -106.3693
17.7515 | -106.3707
17.7552 | -106.3757
17.759 | -106.3807
17.7627 | -106.3853
17.7665 | -106.3934
17.7703 | -106.4031
17.774 | -106.4108
17.7778 | -106.4177
17.7815 | -106.4277
17.7853 | -106.436
17.7891 | -106.4473
17.7928 | -106.4588
17.7966 | -106.468
17.8003 | -106.4739
17.8041 | -106.4801
17.8078 | -106.4872
17.8116 | -106.4964
17.8154 | -106.5047
17.8191 | -106.5099
17.8229 | -106.5122
17.8266 | -106.5151
17.8304 | -106.5159
17.8342 | -106.5145
17.8379 | -106.5076
17.8417 | -106.5012
17.8454 | -106.4939
17.8492 | -106.4857
17.8529 | -106.476
17.8567 | -106.466
17.8605 | -106.4579
17.8642 | -106.4449
17.868 | -106.4275
17.8717 | -106.4111
17.8755 | -106.3981
17.8793 | -106.3843
17.883 | -106.3693
17.8868 | -106.3554
17.8905 | -106.3449
17.8943 | -106.3311
17.8981 | -106.3199
17.9018 | -106.3073
17.9056 | -106.2956
17.9093 | -106.2845
17.9131 | -106.2763
17.9168 | -106.2694
17.9206 | -106.2632
17.9244 | -106.2591
17.9281 | -106.2591
17.9319 | -106.2602
17.9356 | -106.2614
17.9394 | -106.2627
17.9432 | -106.2644
17.9469 | -106.2695
17.9507 | -106.2769
17.9544 | -106.2844
17.9582 | -106.2904
17.9619 | -106.297
17.9657 | -106.3063
17.9695 | -106.3161
17.9732 | -106.3232
17.977 | -106.3289
17.9807 | -106.3366
17.9845 | -106.3477
17.9883 | -106.3611
17.992 | -106.3723
17.9958 | -106.3822
17.9995 | -106.3934
18.0033 | -106.4042
18.007 | -106.4129
18.0108 | -106.4204
18.0146 | -106.4317
18.0183 | -106.4431
18.0221 | -106.4533
18.0258 | -106.464
18.0296 | -106.4776
18.0334 | -106.4898
18.0371 | -106.4974
18.0409 | -106.5063
18.0446 | -106.5152
18.0484 | -106.5255
18.0521 | -106.537
18.0559 | -106.5511
18.0597 | -106.5658
18.0634 | -106.5768
18.0672 | -106.588
18.0709 | -106.6007
18.0747 | -106.6122
18.0785 | -106.6256
18.0822 | -106.6375
18.086 | -106.6512
18.0897 | -106.6658
18.0935 | -106.6839
18.0973 | -106.7002
18.101 | -106.714
18.1048 | -106.7283
18.1085 | -106.7404
18.1123 | -106.7542
18.116 | -106.7679
18.1198 | -106.7823
18.1236 | -106.7974
18.1273 | -106.8094
18.1311 | -106.8207
18.1348 | -106.8317
18.1386 | -106.8447
18.1424 | -106.857
18.1461 | -106.868
18.1499 | -106.8776
18.1536 | -106.886
18.1574 | -106.8956
18.1611 | -106.9068
18.1649 | -106.9155
18.1687 | -106.9216
18.1724 | -106.9262
18.1762 | -106.9309
18.1799 | -106.9329
18.1837 | -106.9333
18.1875 | -106.9321
18.1912 | -106.9312
18.195 | -106.9312
18.1987 | -106.9293
18.2025 | -106.9262
18.2062 | -106.9205
18.21 | -106.9131
18.2138 | -106.9089
18.2175 | -106.9042
18.2213 | -106.8957
18.225 | -106.8845
18.2288 | -106.8744
18.2326 | -106.8661
18.2363 | -106.8588
18.2401 | -106.8471
18.2438 | -106.8337
18.2476 | -106.8194
18.2514 | -106.8048
18.2551 | -106.7925
18.2589 | -106.7824
18.2626 | -106.7715
18.2664 | -106.761
18.2701 | -106.7506
18.2739 | -106.7381
18.2777 | -106.7254
18.2814 | -106.7111
18.2852 | -106.6982
18.2889 | -106.6862
18.2927 | -106.6746
18.2965 | -106.6676
18.3002 | -106.6603
18.304 | -106.6507
18.3077 | -106.6386
18.3115 | -106.6281
18.3152 | -106.6186
18.319 | -106.6111
18.3228 | -106.6024
18.3265 | -106.5922
18.3303 | -106.583
18.334 | -106.577
18.3378 | -106.5716
18.3416 | -106.5668
18.3453 | -106.5634
18.3491 | -106.5612
18.3528 | -106.5589
18.3566 | -106.5563
18.3603 | -106.5555
18.3641 | -106.5571
18.3679 | -106.5593
18.3716 | -106.5595
18.3754 | -106.5634
18.3791 | -106.5656
18.3829 | -106.5644
18.3867 | -106.5658
18.3904 | -106.5665
18.3942 | -106.5677
18.3979 | -106.5686
18.4017 | -106.5672
18.4054 | -106.5661
18.4092 | -106.5681
18.413 | -106.5685
18.4167 | -106.5657
18.4205 | -106.5627
18.4242 | -106.5606
18.428 | -106.5606
18.4318 | -106.5611
18.4355 | -106.56
18.4393 | -106.5584
18.443 | -106.5573
18.4468 | -106.5569
18.4506 | -106.556
18.4543 | -106.5527
18.4581 | -106.5492
18.4618 | -106.5472
18.4656 | -106.5452
18.4693 | -106.5447
18.4731 | -106.5452
18.4769 | -106.5433
18.4806 | -106.5457
18.4844 | -106.5512
18.4881 | -106.554
18.4919 | -106.555
18.4957 | -106.5571
18.4994 | -106.5607
18.5032 | -106.5668
18.5069 | -106.5717
18.5107 | -106.5759
18.5144 | -106.5824
18.5182 | -106.5876
18.522 | -106.5898
18.5257 | -106.5918
18.5295 | -106.5946
18.5332 | -106.5995
18.537 | -106.6038
18.5408 | -106.6077
18.5445 | -106.6129
18.5483 | -106.6174
18.552 | -106.623
18.5558 | -106.6277
18.5595 | -106.6297
18.5633 | -106.629
18.5671 | -106.6297
18.5708 | -106.6318
18.5746 | -106.6341
18.5783 | -106.6323
18.5821 | -106.6272
18.5859 | -106.6241
18.5896 | -106.6231
18.5934 | -106.6175
18.5971 | -106.6117
18.6009 | -106.6075
18.6047 | -106.6009
18.6084 | -106.5935
18.6122 | -106.5862
18.6159 | -106.5786
18.6197 | -106.5735
18.6234 | -106.5689
18.6272 | -106.563
18.631 | -106.5589
18.6347 | -106.5545
18.6385 | -106.5488
18.6422 | -106.5452
18.646 | -106.546
18.6498 | -106.5444
18.6535 | -106.543
18.6573 | -106.5407
18.661 | -106.5395
18.6648 | -106.54
18.6685 | -106.5433
18.6723 | -106.5447
18.6761 | -106.545
18.6798 | -106.5478
18.6836 | -106.5523
18.6873 | -106.5573
18.6911 | -106.5622
18.6949 | -106.5674
18.6986 | -106.5744
18.7024 | -106.5822
18.7061 | -106.5902
18.7099 | -106.5988
18.7136 | -106.6088
18.7174 | -106.6184
18.7212 | -106.6285
18.7249 | -106.6381
18.7287 | -106.6472
18.7324 | -106.6573
18.7362 | -106.6674
18.74 | -106.6747
18.7437 | -106.6809
18.7475 | -106.6865
18.7512 | -106.693
18.755 | -106.7022
18.7587 | -106.7111
18.7625 | -106.7182
18.7663 | -106.7228
18.77 | -106.7273
18.7738 | -106.7312
18.7775 | -106.733
18.7813 | -106.7356
18.7851 | -106.7372
18.7888 | -106.7378
18.7926 | -106.7377
18.7963 | -106.7357
18.8001 | -106.7334
18.8039 | -106.7304
18.8076 | -106.725
18.8114 | -106.7175
18.8151 | -106.7091
18.8189 | -106.6992
18.8226 | -106.6916
18.8264 | -106.683
18.8302 | -106.6769
18.8339 | -106.6697
18.8377 | -106.6604
18.8414 | -106.6466
18.8452 | -106.6363
18.849 | -106.6267
18.8527 | -106.6182
18.8565 | -106.6097
18.8602 | -106.604
18.864 | -106.5947
18.8677 | -106.5841
18.8715 | -106.5744
18.8753 | -106.5652
18.879 | -106.5568
18.8828 | -106.5486
18.8865 | -106.5419
18.8903 | -106.5347
18.8941 | -106.5265
18.8978 | -106.5183
18.9016 | -106.5107
18.9053 | -106.5039
18.9091 | -106.4978
18.9128 | -106.4912
18.9166 | -106.487
18.9204 | -106.4818
18.9241 | -106.4761
18.9279 | -106.4725
18.9316 | -106.4697
18.9354 | -106.4639
18.9392 | -106.4597
18.9429 | -106.4567
18.9467 | -106.4514
18.9504 | -106.4462
18.9542 | -106.442
18.958 | -106.4376
18.9617 | -106.4318
18.9655 | -106.4272
18.9692 | -106.4222
18.973 | -106.4168
18.9767 | -106.4072
18.9805 | -106.3992
18.9843 | -106.3909
18.988 | -106.3817
18.9918 | -106.375
18.9955 | -106.3667
18.9993 | -106.3588
19.0031 | -106.3493
19.0068 | -106.3387
19.0106 | -106.3299
19.0143 | -106.3208
19.0181 | -106.3111
19.0218 | -106.2994
19.0256 | -106.2875
19.0294 | -106.2777
19.0331 | -106.2663
19.0369 | -106.2555
19.0406 | -106.244
19.0444 | -106.2324
19.0482 | -106.2225
19.0519 | -106.2125
19.0557 | -106.2046
19.0594 | -106.1963
19.0632 | -106.1887
19.0669 | -106.1825
19.0707 | -106.175
19.0745 | -106.1677
19.0782 | -106.1645
19.082 | -106.1611
19.0857 | -106.1575
19.0895 | -106.155
19.0933 | -106.1528
19.097 | -106.1521
19.1008 | -106.1511
19.1045 | -106.1503
19.1083 | -106.1534
19.1121 | -106.1584
19.1158 | -106.1637
19.1196 | -106.169
19.1233 | -106.1716
19.1271 | -106.1734
19.1308 | -106.1769
19.1346 | -106.1826
19.1384 | -106.1883
19.1421 | -106.1955
19.1459 | -106.203
19.1496 | -106.2109
19.1534 | -106.2171
19.1572 | -106.2219
19.1609 | -106.2288
19.1647 | -106.2374
19.1684 | -106.2477
19.1722 | -106.2598
19.1759 | -106.2689
19.1797 | -106.2775
19.1835 | -106.2861
19.1872 | -106.2967
19.191 | -106.3076
19.1947 | -106.3184
19.1985 | -106.3282
19.2023 | -106.3375
19.206 | -106.3466
19.2098 | -106.3546
19.2135 | -106.3613
19.2173 | -106.3685
19.221 | -106.3747
19.2248 | -106.3778
19.2286 | -106.3815
19.2323 | -106.3842
19.2361 | -106.3862
19.2398 | -106.3871
19.2436 | -106.3867
19.2474 | -106.3828
19.2511 | -106.382
19.2549 | -106.3809
19.2586 | -106.3756
19.2624 | -106.3706
19.2661 | -106.3643
19.2699 | -106.357
19.2737 | -106.3479
19.2774 | -106.3371
19.2812 | -106.3256
19.2849 | -106.3122
19.2887 | -106.2998
19.2925 | -106.2875
19.2962 | -106.2752
19.3 | -106.2633
19.3037 | -106.251
19.3075 | -106.2372
19.3113 | -106.2252
19.315 | -106.2138
19.3188 | -106.2026
19.3225 | -106.19
19.3263 | -106.1766
19.33 | -106.1625
19.3338 | -106.1516
19.3376 | -106.1417
19.3413 | -106.1293
19.3451 | -106.1203
19.3488 | -106.1128
19.3526 | -106.1037
19.3564 | -106.0971
19.3601 | -106.0931
19.3639 | -106.0884
19.3676 | -106.0837
19.3714 | -106.0808
19.3751 | -106.0816
19.3789 | -106.083
19.3827 | -106.0836
19.3864 | -106.0859
19.3902 | -106.0887
19.3939 | -106.094
19.3977 | -106.1028
19.4015 | -106.1126
19.4052 | -106.1223
19.409 | -106.1329
19.4127 | -106.142
19.4165 | -106.1535
19.4202 | -106.1647
19.424 | -106.1758
19.4278 | -106.1877
19.4315 | -106.2006
19.4353 | -106.2133
19.439 | -106.2263
19.4428 | -106.239
19.4466 | -106.2479
19.4503 | -106.2572
19.4541 | -106.2681
19.4578 | -106.2773
19.4616 | -106.2836
19.4654 | -106.2901
19.4691 | -106.2967
19.4729 | -106.3021
19.4766 | -106.3077
19.4804 | -106.3129
19.4841 | -106.3202
19.4879 | -106.3264
19.4917 | -106.3304
19.4954 | -106.3329
19.4992 | -106.3341
19.5029 | -106.3337
19.5067 | -106.335
19.5105 | -106.3345
19.5142 | -106.3332
19.518 | -106.3284
19.5217 | -106.3291
19.5255 | -106.3302
19.5292 | -106.3273
19.533 | -106.3232
19.5368 | -106.3215
19.5405 | -106.3193
19.5443 | -106.3141
19.548 | -106.3098
19.5518 | -106.3069
19.5556 | -106.3052
19.5593 | -106.3052
19.5631 | -106.3037
19.5668 | -106.301
19.5706 | -106.2974
19.5743 | -106.2941
19.5781 | -106.2878
19.5819 | -106.2813
19.5856 | -106.2791
19.5894 | -106.2739
19.5931 | -106.2684
19.5969 | -106.2632
19.6007 | -106.26
19.6044 | -106.2556
19.6082 | -106.2516
19.6119 | -106.2469
19.6157 | -106.2405
19.6195 | -106.2363
19.6232 | -106.2311
19.627 | -106.227
19.6307 | -106.2225
19.6345 | -106.2159
19.6382 | -106.2089
19.642 | -106.2036
19.6458 | -106.1993
19.6495 | -106.1973
19.6533 | -106.1932
19.657 | -106.1891
19.6608 | -106.1842
19.6646 | -106.1786
19.6683 | -106.1735
19.6721 | -106.1664
19.6758 | -106.1603
19.6796 | -106.1531
19.6833 | -106.1466
19.6871 | -106.1417
19.6909 | -106.1381
19.6946 | -106.1307
19.6984 | -106.1227
19.7021 | -106.1146
19.7059 | -106.1085
19.7097 | -106.1009
19.7134 | -106.0925
19.7172 | -106.0822
19.7209 | -106.0713
19.7247 | -106.062
19.7284 | -106.053
19.7322 | -106.0418
19.736 | -106.0322
19.7397 | -106.0218
19.7435 | -106.0111
19.7472 | -106.0007
19.751 | -105.9885
19.7548 | -105.9761
19.7585 | -105.9636
19.7623 | -105.9496
19.766 | -105.9355
19.7698 | -105.9229
19.7735 | -105.9098
19.7773 | -105.8984
19.7811 | -105.885
19.7848 | -105.8711
19.7886 | -105.8585
19.7923 | -105.8457
19.7961 | -105.8335
19.7999 | -105.8221
19.8036 | -105.8117
19.8074 | -105.8016
19.8111 | -105.7922
19.8149 | -105.7836
19.8187 | -105.7753
19.8224 | -105.7689
19.8262 | -105.7639
19.8299 | -105.7583
19.8337 | -105.7549
19.8374 | -105.751
19.8412 | -105.745
19.845 | -105.7434
19.8487 | -105.7409
19.8525 | -105.7402
19.8562 | -105.7402
19.86 | -105.7418
19.8638 | -105.744
19.8675 | -105.7433
19.8713 | -105.744
19.875 | -105.7463
19.8788 | -105.7508
19.8825 | -105.7534
19.8863 | -105.7548
19.8901 | -105.7558
19.8938 | -105.7538
19.8976 | -105.7529
19.9013 | -105.7499
19.9051 | -105.743
19.9089 | -105.735
19.9126 | -105.728
19.9164 | -105.7218
19.9201 | -105.7158
19.9239 | -105.7096
19.9276 | -105.7001
19.9314 | -105.6881
19.9352 | -105.6735
19.9389 | -105.66
19.9427 | -105.6481
19.9464 | -105.6363
19.9502 | -105.6226
19.954 | -105.6072
19.9577 | -105.5925
19.9615 | -105.5795
19.9652 | -105.5647
19.969 | -105.5478
19.9728 | -105.5323
19.9765 | -105.5172
19.9803 | -105.5044
19.984 | -105.4919
19.9878 | -105.4776
19.9915 | -105.4647
19.9953 | -105.452
19.9991 | -105.4387
20.0028 | -105.4258
20.0066 | -105.4167
20.0103 | -105.4088
20.0141 | -105.4008
20.0179 | -105.3933
20.0216 | -105.3871
20.0254 | -105.3825
20.0291 | -105.3791
20.0329 | -105.3755
20.0366 | -105.3728
20.0404 | -105.3705
20.0442 | -105.3713
20.0479 | -105.3721
20.0517 | -105.3723
20.0554 | -105.3756
20.0592 | -105.3814
20.063 | -105.384
20.0667 | -105.3878
20.0705 | -105.3898
20.0742 | -105.3919
20.078 | -105.3954
20.0817 | -105.3989
20.0855 | -105.4022
20.0893 | -105.4046
20.093 | -105.4088
20.0968 | -105.4137
20.1005 | -105.4196
20.1043 | -105.4241
20.1081 | -105.4277
20.1118 | -105.4306
20.1156 | -105.4346
20.1193 | -105.4375
20.1231 | -105.4419
20.1269 | -105.4492
20.1306 | -105.4563
20.1344 | -105.4635
20.1381 | -105.4688
20.1419 | -105.4743
20.1456 | -105.4802
20.1494 | -105.4883
20.1532 | -105.4991
20.1569 | -105.5081
20.1607 | -105.5152
20.1644 | -105.5197
20.1682 | -105.5258
20.172 | -105.534
20.1757 | -105.5438
20.1795 | -105.5533
20.1832 | -105.5628
20.187 | -105.5735
20.1907 | -105.5836
20.1945 | -105.5934
20.1983 | -105.6024
20.202 | -105.6104
20.2058 | -105.6162
20.2095 | -105.6195
20.2133 | -105.6241
20.2171 | -105.6294
20.2208 | -105.6362
20.2246 | -105.6422
20.2283 | -105.6476
20.2321 | -105.6513
20.2358 | -105.6528
20.2396 | -105.656
20.2434 | -105.6579
20.2471 | -105.6574
20.2509 | -105.6582
20.2546 | -105.66
20.2584 | -105.6615
20.2622 | -105.6627
20.2659 | -105.6622
20.2697 | -105.6638
20.2734 | -105.6651
20.2772 | -105.6642
20.2809 | -105.6643
20.2847 | -105.6657
20.2885 | -105.6655
20.2922 | -105.6641
20.296 | -105.6631
20.2997 | -105.6618
20.3035 | -105.6604
20.3073 | -105.6593
20.311 | -105.6597
20.3148 | -105.6605
20.3185 | -105.6622
20.3223 | -105.6644
20.3261 | -105.6681
20.3298 | -105.6693
20.3336 | -105.6712
20.3373 | -105.6736
20.3411 | -105.6764
20.3448 | -105.6809
20.3486 | -105.6846
20.3524 | -105.6887
20.3561 | -105.6921
20.3599 | -105.6974
20.3636 | -105.7027
20.3674 | -105.707
20.3712 | -105.7099
20.3749 | -105.7138
20.3787 | -105.7188
20.3824 | -105.723
20.3862 | -105.728
20.3899 | -105.7322
20.3937 | -105.7365
20.3975 | -105.7423
20.4012 | -105.7492
20.405 | -105.7551
20.4087 | -105.7608
20.4125 | -105.7655
20.4163 | -105.7701
20.42 | -105.7739
20.4238 | -105.7792
20.4275 | -105.7848
20.4313 | -105.7892
20.435 | -105.7951
20.4388 | -105.7991
20.4426 | -105.8023
20.4463 | -105.8076
20.4501 | -105.8136
20.4538 | -105.8196
20.4576 | -105.824
20.4614 | -105.8269
20.4651 | -105.8318
20.4689 | -105.8383
20.4726 | -105.8421
20.4764 | -105.8453
20.4801 | -105.8512
20.4839 | -105.8563
20.4877 | -105.8617
20.4914 | -105.8673
20.4952 | -105.8724
20.4989 | -105.8781
20.5027 | -105.884
20.5065 | -105.8895
20.5102 | -105.8958
20.514 | -105.8997
20.5177 | -105.9021
20.5215 | -105.9068
20.5253 | -105.9127
20.529 | -105.9172
20.5328 | -105.9219
20.5365 | -105.9265
20.5403 | -105.9289
20.544 | -105.9313
20.5478 | -105.9336
20.5516 | -105.9364
20.5553 | -105.9395
20.5591 | -105.9413
20.5628 | -105.9451
20.5666 | -105.9468
20.5704 | -105.9458
20.5741 | -105.9446
20.5779 | -105.9443
20.5816 | -105.9429
20.5854 | -105.9383
20.5891 | -105.9355
20.5929 | -105.9318
20.5967 | -105.9259
20.6004 | -105.9147
20.6042 | -105.8965
20.6079 | -105.886
20.6117 | -105.8756
20.6155 | -105.8621
20.6192 | -105.8466
20.623 | -105.8313
20.6267 | -105.8151
20.6305 | -105.7993
20.6342 | -105.7827
20.638 | -105.763
20.6418 | -105.7434
20.6455 | -105.7237
20.6493 | -105.7049
20.653 | -105.6846
20.6568 | -105.6646
20.6606 | -105.6438
20.6643 | -105.6239
20.6681 | -105.6026
20.6718 | -105.5787
20.6756 | -105.554
20.6794 | -105.5288
20.6831 | -105.5046
20.6869 | -105.4817
20.6906 | -105.4579
20.6944 | -105.4365
20.6981 | -105.4164
20.7019 | -105.3956
20.7057 | -105.3754
20.7094 | -105.3542
20.7132 | -105.3345
20.7169 | -105.3166
20.7207 | -105.2983
20.7245 | -105.2824
20.7282 | -105.2674
20.732 | -105.2528
20.7357 | -105.2403
20.7395 | -105.2284
20.7432 | -105.2183
20.747 | -105.2084
20.7508 | -105.2016
20.7545 | -105.1961
20.7583 | -105.1895
20.762 | -105.1837
20.7658 | -105.1803
20.7696 | -105.1771
20.7733 | -105.1735
20.7771 | -105.1722
20.7808 | -105.1719
20.7846 | -105.1722
20.7883 | -105.1725
20.7921 | -105.1759
20.7959 | -105.1802
20.7996 | -105.1842
20.8034 | -105.188
20.8071 | -105.1939
20.8109 | -105.1983
20.8147 | -105.2016
20.8184 | -105.2055
20.8222 | -105.2097
20.8259 | -105.2137
20.8297 | -105.218
20.8335 | -105.2227
20.8372 | -105.2264
20.841 | -105.23
20.8447 | -105.2323
20.8485 | -105.233
20.8522 | -105.2324
20.856 | -105.2313
20.8598 | -105.2319
20.8635 | -105.2314
20.8673 | -105.2306
20.871 | -105.2286
20.8748 | -105.2247
20.8786 | -105.221
20.8823 | -105.2144
20.8861 | -105.2079
20.8898 | -105.2012
20.8936 | -105.197
20.8973 | -105.1908
20.9011 | -105.1812
20.9049 | -105.1718
20.9086 | -105.1627
20.9124 | -105.152
20.9161 | -105.1446
20.9199 | -105.138
20.9237 | -105.1319
20.9274 | -105.1238
20.9312 | -105.1157
20.9349 | -105.1096
20.9387 | -105.1043
20.9424 | -105.1004
20.9462 | -105.0963
20.95 | -105.0928
20.9537 | -105.0892
20.9575 | -105.088
20.9612 | -105.0879
20.965 | -105.0886
20.9688 | -105.0908
20.9725 | -105.0965
20.9763 | -105.1033
20.98 | -105.1101
20.9838 | -105.1165
20.9875 | -105.1252
20.9913 | -105.1324
20.9951 | -105.1411
20.9988 | -105.1517
21.0026 | -105.1631
21.0063 | -105.1766
21.0101 | -105.1892
21.0139 | -105.2032
21.0176 | -105.2196
21.0214 | -105.236
21.0251 | -105.253
21.0289 | -105.2694
21.0327 | -105.2841
21.0364 | -105.2995
21.0402 | -105.3154
21.0439 | -105.3318
21.0477 | -105.3461
21.0514 | -105.3627
21.0552 | -105.3806
21.059 | -105.3984
21.0627 | -105.4162
21.0665 | -105.4343
21.0702 | -105.4502
21.074 | -105.4619
21.0778 | -105.4733
21.0815 | -105.4862
21.0853 | -105.502
21.089 | -105.515
21.0928 | -105.5268
21.0965 | -105.5369
21.1003 | -105.5452
21.1041 | -105.5518
21.1078 | -105.5601
21.1116 | -105.5694
21.1153 | -105.5771
21.1191 | -105.5816
21.1229 | -105.5815
21.1266 | -105.5819
21.1304 | -105.583
21.1341 | -105.5855
21.1379 | -105.5842
21.1416 | -105.5835
21.1454 | -105.5777
21.1492 | -105.5742
21.1529 | -105.5715
21.1567 | -105.5696
21.1604 | -105.5664
21.1642 | -105.5618
21.168 | -105.5544
21.1717 | -105.5478
21.1755 | -105.5391
21.1792 | -105.5316
21.183 | -105.5255
21.1868 | -105.5173
21.1905 | -105.5088
21.1943 | -105.4998
21.198 | -105.4918
21.2018 | -105.4844
21.2055 | -105.4744
21.2093 | -105.4639
21.2131 | -105.4573
21.2168 | -105.4477
21.2206 | -105.436
21.2243 | -105.4256
21.2281 | -105.4178
21.2319 | -105.4097
21.2356 | -105.4036
21.2394 | -105.3987
21.2431 | -105.3913
21.2469 | -105.3831
21.2506 | -105.3772
21.2544 | -105.3728
21.2582 | -105.369
21.2619 | -105.3659
21.2657 | -105.3628
21.2694 | -105.3598
21.2732 | -105.356
21.277 | -105.3517
21.2807 | -105.3497
21.2845 | -105.3502
21.2882 | -105.3512
21.292 | -105.3528
21.2957 | -105.353
21.2995 | -105.3553
21.3033 | -105.3599
21.307 | -105.3633
21.3108 | -105.3663
21.3145 | -105.3699
21.3183 | -105.3738
21.3221 | -105.3794
21.3258 | -105.3857
21.3296 | -105.3949
21.3333 | -105.4047
21.3371 | -105.4147
21.3408 | -105.424
21.3446 | -105.4352
21.3484 | -105.4479
21.3521 | -105.4576
21.3559 | -105.4648
21.3596 | -105.4728
21.3634 | -105.4808
21.3672 | -105.4899
21.3709 | -105.5002
21.3747 | -105.5112
21.3784 | -105.5232
21.3822 | -105.5325
21.386 | -105.5387
21.3897 | -105.5451
21.3935 | -105.5529
21.3972 | -105.5621
21.401 | -105.5704
21.4047 | -105.5795
21.4085 | -105.586
21.4123 | -105.5888
21.416 | -105.5901
21.4198 | -105.5933
21.4235 | -105.5979
21.4273 | -105.5987
21.4311 | -105.5986
21.4348 | -105.5948
21.4386 | -105.593
21.4423 | -105.5897
21.4461 | -105.586
21.4498 | -105.5826
21.4536 | -105.5788
21.4574 | -105.5739
21.4611 | -105.5669
21.4649 | -105.5596
21.4686 | -105.5522
21.4724 | -105.5464
21.4762 | -105.5388
21.4799 | -105.5306
21.4837 | -105.5224
21.4874 | -105.5136
21.4912 | -105.5053
21.4949 | -105.4983
21.4987 | -105.4903
21.5025 | -105.4829
21.5062 | -105.4758
21.51 | -105.4679
21.5137 | -105.4601
21.5175 | -105.4514
21.5213 | -105.4435
21.525 | -105.437
21.5288 | -105.4301
21.5325 | -105.4229
21.5363 | -105.4147
21.5401 | -105.4067
21.5438 | -105.401
21.5476 | -105.3965
21.5513 | -105.3934
21.5551 | -105.3905
21.5588 | -105.387
21.5626 | -105.383
21.5664 | -105.3794
21.5701 | -105.3779
21.5739 | -105.3769
21.5776 | -105.3743
21.5814 | -105.3746
21.5852 | -105.374
21.5889 | -105.3717
21.5927 | -105.3687
21.5964 | -105.3656
21.6002 | -105.3646
21.6039 | -105.3645
21.6077 | -105.3644
21.6115 | -105.3647
21.6152 | -105.3662
21.619 | -105.3662
21.6227 | -105.3645
21.6265 | -105.3627
21.6303 | -105.3612
21.634 | -105.3594
21.6378 | -105.3604
21.6415 | -105.3619
21.6453 | -105.3618
21.649 | -105.3602
21.6528 | -105.3572
21.6566 | -105.3564
21.6603 | -105.3564
21.6641 | -105.357
21.6678 | -105.3549
21.6716 | -105.3499
21.6754 | -105.3461
21.6791 | -105.3436
21.6829 | -105.3387
21.6866 | -105.3304
21.6904 | -105.3242
21.6942 | -105.3173
21.6979 | -105.3097
21.7017 | -105.304
21.7054 | -105.2984
21.7092 | -105.2933
21.7129 | -105.2858
21.7167 | -105.278
21.7205 | -105.2716
21.7242 | -105.2663
21.728 | -105.2588
21.7317 | -105.2502
21.7355 | -105.2438
21.7393 | -105.237
21.743 | -105.2309
21.7468 | -105.2249
21.7505 | -105.2204
21.7543 | -105.2161
21.758 | -105.2111
21.7618 | -105.2054
21.7656 | -105.1987
21.7693 | -105.1937
21.7731 | -105.1894
21.7768 | -105.1856
21.7806 | -105.181
21.7844 | -105.1778
21.7881 | -105.1765
21.7919 | -105.1745
21.7956 | -105.1722
21.7994 | -105.1702
21.8031 | -105.1675
21.8069 | -105.1666
21.8107 | -105.1631
21.8144 | -105.1615
21.8182 | -105.1629
21.8219 | -105.165
21.8257 | -105.1664
21.8295 | -105.1678
21.8332 | -105.1697
21.837 | -105.1732
21.8407 | -105.1771
21.8445 | -105.1821
21.8483 | -105.1855
21.852 | -105.1867
21.8558 | -105.1906
21.8595 | -105.1937
21.8633 | -105.1943
21.867 | -105.1964
21.8708 | -105.2
21.8746 | -105.2026
21.8783 | -105.2039
21.8821 | -105.2066
21.8858 | -105.2106
21.8896 | -105.2134
21.8934 | -105.2145
21.8971 | -105.219
21.9009 | -105.223
21.9046 | -105.2285
21.9084 | -105.2327
21.9121 | -105.2345
21.9159 | -105.2366
21.9197 | -105.2394
21.9234 | -105.2435
21.9272 | -105.248
21.9309 | -105.2536
21.9347 | -105.2584
21.9385 | -105.2619
21.9422 | -105.2659
21.946 | -105.2699
21.9497 | -105.2732
21.9535 | -105.2795
21.9572 | -105.285
21.961 | -105.2911
21.9648 | -105.295
21.9685 | -105.2982
21.9723 | -105.3046
21.976 | -105.313
21.9798 | -105.3225
21.9836 | -105.3283
21.9873 | -105.3346
21.9911 | -105.3424
21.9948 | -105.35
21.9986 | -105.3583
22.0023 | -105.3668
22.0061 | -105.3758
22.0099 | -105.3841
22.0136 | -105.3872
22.0174 | -105.3897
22.0211 | -105.3927
22.0249 | -105.396
22.0287 | -105.4009
22.0324 | -105.4059
22.0362 | -105.4112
22.0399 | -105.4148
22.0437 | -105.4178
22.0475 | -105.4193
22.0512 | -105.4198
22.055 | -105.4209
22.0587 | -105.4225
22.0625 | -105.4219
22.0662 | -105.4213
22.07 | -105.4213
22.0738 | -105.4216
22.0775 | -105.4215
22.0813 | -105.4199
22.085 | -105.4185
22.0888 | -105.4175
22.0926 | -105.4187
22.0963 | -105.4178
22.1001 | -105.4168
22.1038 | -105.4159
22.1076 | -105.4153
22.1113 | -105.415
22.1151 | -105.4151
22.1189 | -105.4155
22.1226 | -105.4184
22.1264 | -105.4199
22.1301 | -105.4205
22.1339 | -105.4206
22.1377 | -105.4206
22.1414 | -105.4212
22.1452 | -105.4218
22.1489 | -105.4231
22.1527 | -105.4242
22.1564 | -105.4259
22.1602 | -105.4282
22.164 | -105.4297
22.1677 | -105.4321
22.1715 | -105.4358
22.1752 | -105.4396
22.179 | -105.4431
22.1828 | -105.4459
22.1865 | -105.4484
22.1903 | -105.4526
22.194 | -105.4566
22.1978 | -105.4607
22.2016 | -105.4631
22.2053 | -105.4665
22.2091 | -105.4662
22.2128 | -105.4663
22.2166 | -105.4667
22.2203 | -105.4679
22.2241 | -105.4696
22.2279 | -105.4712
22.2316 | -105.4683
22.2354 | -105.4693
22.2391 | -105.4691
22.2429 | -105.4685
22.2467 | -105.4682
22.2504 | -105.4678
22.2542 | -105.4682
22.2579 | -105.4659
22.2617 | -105.4633
22.2654 | -105.4606
22.2692 | -105.4569
22.273 | -105.4544
22.2767 | -105.4516
22.2805 | -105.4491
22.2842 | -105.4445
22.288 | -105.4366
22.2918 | -105.4293
22.2955 | -105.4213
22.2993 | -105.4155
22.303 | -105.4094
22.3068 | -105.4014
22.3105 | -105.3938
22.3143 | -105.3864
22.3181 | -105.3774
22.3218 | -105.3711
22.3256 | -105.3658
22.3293 | -105.3604
22.3331 | -105.3532
22.3369 | -105.3479
22.3406 | -105.3444
22.3444 | -105.3389
22.3481 | -105.3314
22.3519 | -105.3247
22.3557 | -105.322
22.3594 | -105.3198
22.3632 | -105.3171
22.3669 | -105.3149
22.3707 | -105.3121
22.3744 | -105.3088
22.3782 | -105.3071
22.382 | -105.3056
22.3857 | -105.3068
22.3895 | -105.3061
22.3932 | -105.3053
22.397 | -105.3046
22.4008 | -105.3023
22.4045 | -105.3026
22.4083 | -105.3025
22.412 | -105.3036
22.4158 | -105.3055
22.4195 | -105.3078
22.4233 | -105.3114
22.4271 | -105.3141
22.4308 | -105.3169
22.4346 | -105.3198
22.4383 | -105.3236
22.4421 | -105.3279
22.4459 | -105.3327
22.4496 | -105.3363
22.4534 | -105.3394
22.4571 | -105.3418
22.4609 | -105.345
22.4646 | -105.3469
22.4684 | -105.3497
22.4722 | -105.3521
22.4759 | -105.356
22.4797 | -105.3591
22.4834 | -105.3605
22.4872 | -105.362
22.491 | -105.3646
22.4947 | -105.3649
22.4985 | -105.3643
22.5022 | -105.3648
22.506 | -105.365
22.5097 | -105.3639
22.5135 | -105.3612
22.5173 | -105.3592
22.521 | -105.3572
22.5248 | -105.3559
22.5285 | -105.3525
22.5323 | -105.3468
22.5361 | -105.3413
22.5398 | -105.3368
22.5436 | -105.3313
22.5473 | -105.3251
22.5511 | -105.318
22.5549 | -105.3089
22.5586 | -105.3008
22.5624 | -105.2931
22.5661 | -105.2842
22.5699 | -105.2759
22.5736 | -105.2661
22.5774 | -105.2595
22.5812 | -105.2515
22.5849 | -105.2451
22.5887 | -105.2389
22.5924 | -105.2323
22.5962 | -105.224
22.6 | -105.2176
22.6037 | -105.2146
22.6075 | -105.21
22.6112 | -105.2061
22.615 | -105.2024
22.6187 | -105.1979
22.6225 | -105.1936
22.6263 | -105.1901
22.63 | -105.188
22.6338 | -105.1875
22.6375 | -105.1884
22.6413 | -105.187
22.6451 | -105.1855
22.6488 | -105.186
22.6526 | -105.1895
22.6563 | -105.1888
22.6601 | -105.1874
22.6638 | -105.1888
22.6676 | -105.1931
22.6714 | -105.1977
22.6751 | -105.2004
22.6789 | -105.2037
22.6826 | -105.2083
22.6864 | -105.2145
22.6902 | -105.2199
22.6939 | -105.2245
22.6977 | -105.2305
22.7014 | -105.2382
22.7052 | -105.2462
22.709 | -105.2564
22.7127 | -105.2662
22.7165 | -105.2748
22.7202 | -105.2859
22.724 | -105.2938
22.7277 | -105.2992
22.7315 | -105.3044
22.7353 | -105.3111
22.739 | -105.3177
22.7428 | -105.3226
22.7465 | -105.3296
22.7503 | -105.3361
22.7541 | -105.3415
22.7578 | -105.3468
22.7616 | -105.3531
22.7653 | -105.3573
22.7691 | -105.36
22.7728 | -105.3626
22.7766 | -105.3712
22.7804 | -105.3784
22.7841 | -105.3829
22.7879 | -105.3864
22.7916 | -105.3902
22.7954 | -105.3942
22.7992 | -105.3991
22.8029 | -105.4045
22.8067 | -105.4084
22.8104 | -105.4123
22.8142 | -105.4169
22.8179 | -105.423
22.8217 | -105.4284
22.8255 | -105.433
22.8292 | -105.437
22.833 | -105.4426
22.8367 | -105.4492
22.8405 | -105.4569
22.8443 | -105.4649
22.848 | -105.4731
22.8518 | -105.4823
22.8555 | -105.4925
22.8593 | -105.5028
22.863 | -105.5117
22.8668 | -105.5201
22.8706 | -105.5294
22.8743 | -105.5409
22.8781 | -105.5504
22.8818 | -105.5581
22.8856 | -105.5657
22.8894 | -105.5747
22.8931 | -105.5834
22.8969 | -105.5922
22.9006 | -105.5988
22.9044 | -105.6071
22.9082 | -105.6143
22.9119 | -105.622
22.9157 | -105.6292
22.9194 | -105.6369
22.9232 | -105.6425
22.9269 | -105.646
22.9307 | -105.6501
22.9345 | -105.653
22.9382 | -105.6571
22.942 | -105.6607
22.9457 | -105.6634
22.9495 | -105.6662
22.9533 | -105.6673
22.957 | -105.6675
22.9608 | -105.6673
22.9645 | -105.6666
22.9683 | -105.6655
22.972 | -105.6645
22.9758 | -105.6634
22.9796 | -105.6643
22.9833 | -105.6642
22.9871 | -105.6608
22.9908 | -105.655
22.9946 | -105.6488
22.9984 | -105.6433
23.0021 | -105.637
23.0059 | -105.6308
23.0096 | -105.6256
23.0134 | -105.6214
23.0171 | -105.6187
23.0209 | -105.6134
23.0247 | -105.608
23.0284 | -105.6021
23.0322 | -105.5964
23.0359 | -105.5932
23.0397 | -105.5911
23.0435 | -105.5884
23.0472 | -105.5848
23.051 | -105.5843
23.0547 | -105.5849
23.0585 | -105.5838
23.0622 | -105.5839
23.066 | -105.5867
23.0698 | -105.589
23.0735 | -105.5912
23.0773 | -105.5922
23.081 | -105.5948
23.0848 | -105.597
23.0886 | -105.6029
23.0923 | -105.6112
23.0961 | -105.6167
23.0998 | -105.6192
23.1036 | -105.623
23.1074 | -105.6288
23.1111 | -105.6367
23.1149 | -105.645
23.1186 | -105.652
23.1224 | -105.6578
23.1261 | -105.663
23.1299 | -105.6694
23.1337 | -105.6772
23.1374 | -105.6836
23.1412 | -105.6873
23.1449 | -105.6925
23.1487 | -105.699
23.1525 | -105.7077
23.1562 | -105.7143
23.16 | -105.7211
23.1637 | -105.7288
23.1675 | -105.7356
23.1712 | -105.7412
23.175 | -105.7471
23.1788 | -105.7542
23.1825 | -105.7604
23.1863 | -105.7645
23.19 | -105.7695
23.1938 | -105.7762
23.1976 | -105.7835
23.2013 | -105.7883
23.2051 | -105.7962
23.2088 | -105.8057
23.2126 | -105.8138
23.2163 | -105.8191
23.2201 | -105.8229
23.2239 | -105.8266
23.2276 | -105.8305
23.2314 | -105.8346
23.2351 | -105.8385
23.2389 | -105.84
23.2427 | -105.8419
23.2464 | -105.8444
23.2502 | -105.8468
23.2539 | -105.8506
23.2577 | -105.8541
23.2615 | -105.8569
23.2652 | -105.858
23.269 | -105.8595
23.2727 | -105.86
23.2765 | -105.8599
23.2802 | -105.8606
23.284 | -105.8596
23.2878 | -105.8596
23.2915 | -105.8596
23.2953 | -105.86
23.299 | -105.8591
23.3028 | -105.857
23.3066 | -105.8552
23.3103 | -105.8539
23.3141 | -105.8475
23.3178 | -105.8469
23.3216 | -105.847
23.3253 | -105.8451
23.3291 | -105.8437
23.3329 | -105.8422
23.3366 | -105.8385
23.3404 | -105.8334
23.3441 | -105.8282
23.3479 | -105.8246
23.3517 | -105.8233
23.3554 | -105.8198
23.3592 | -105.8159
23.3629 | -105.8134
23.3667 | -105.8096
23.3704 | -105.8066
23.3742 | -105.8018
23.378 | -105.8001
23.3817 | -105.7969
23.3855 | -105.7919
23.3892 | -105.7878
23.393 | -105.7817
23.3968 | -105.7761
23.4005 | -105.7718
23.4043 | -105.7668
23.408 | -105.7633
23.4118 | -105.7594
23.4156 | -105.7571
23.4193 | -105.753
23.4231 | -105.7478
23.4268 | -105.7436
23.4306 | -105.739
23.4343 | -105.7339
23.4381 | -105.7298
23.4419 | -105.7249
23.4456 | -105.7205
23.4494 | -105.7168
23.4531 | -105.7137
23.4569 | -105.7095
23.4607 | -105.7072
23.4644 | -105.705
23.4682 | -105.703
23.4719 | -105.7013
23.4757 | -105.6982
23.4794 | -105.6962
23.4832 | -105.6932
23.487 | -105.6916
23.4907 | -105.6892
23.4945 | -105.6863
23.4982 | -105.6846
23.502 | -105.6851
23.5058 | -105.6852
23.5095 | -105.684
23.5133 | -105.6838
23.517 | -105.684
23.5208 | -105.6844
23.5245 | -105.6849
23.5283 | -105.6857
23.5321 | -105.6869
23.5358 | -105.6855
23.5396 | -105.6837
23.5433 | -105.684
23.5471 | -105.6861
23.5509 | -105.6884
23.5546 | -105.6913
23.5584 | -105.6945
23.5621 | -105.6957
23.5659 | -105.699
23.5696 | -105.7037
23.5734 | -105.7083
23.5772 | -105.7121
23.5809 | -105.715
23.5847 | -105.7182
23.5884 | -105.7227
23.5922 | -105.7273
23.596 | -105.7329
23.5997 | -105.739
23.6035 | -105.7455
23.6072 | -105.7512
23.611 | -105.7575
23.6148 | -105.7647
23.6185 | -105.7715
23.6223 | -105.7791
23.626 | -105.785
23.6298 | -105.7916
23.6335 | -105.7974
23.6373 | -105.8014
23.6411 | -105.807
23.6448 | -105.8154
23.6486 | -105.824
23.6523 | -105.8294
23.6561 | -105.8339
23.6599 | -105.8403
23.6636 | -105.8436
23.6674 | -105.8486
23.6711 | -105.8552
23.6749 | -105.8603
23.6786 | -105.8662
23.6824 | -105.869
23.6862 | -105.8694
23.6899 | -105.8696
23.6937 | -105.8698
23.6974 | -105.8681
23.7012 | -105.8677
23.705 | -105.867
23.7087 | -105.8682
23.7125 | -105.8673
23.7162 | -105.8657
23.72 | -105.864
23.7237 | -105.8627
23.7275 | -105.8633
23.7313 | -105.861
23.735 | -105.8585
23.7388 | -105.8565
23.7425 | -105.8536
23.7463 | -105.8491
23.7501 | -105.8438
23.7538 | -105.8392
23.7576 | -105.8365
23.7613 | -105.8345
23.7651 | -105.834
23.7689 | -105.833
23.7726 | -105.8324
23.7764 | -105.8326
23.7801 | -105.8319
23.7839 | -105.8307
23.7876 | -105.83
23.7914 | -105.8308
23.7952 | -105.8296
23.7989 | -105.8281
23.8027 | -105.829
23.8064 | -105.831
23.8102 | -105.833
23.814 | -105.8345
23.8177 | -105.8369
23.8215 | -105.8417
23.8252 | -105.8464
23.829 | -105.851
23.8327 | -105.8567
23.8365 | -105.8627
23.8403 | -105.8698
23.844 | -105.8779
23.8478 | -105.8858
23.8515 | -105.8944
23.8553 | -105.902
23.8591 | -105.9102
23.8628 | -105.9174
23.8666 | -105.9255
23.8703 | -105.9349
23.8741 | -105.947
23.8778 | -105.9572
23.8816 | -105.9654
23.8854 | -105.9723
23.8891 | -105.9797
23.8929 | -105.9874
23.8966 | -105.9944
23.9004 | -106.0006
23.9042 | -106.0056
23.9079 | -106.0088
23.9117 | -106.0121
23.9154 | -106.0172
23.9192 | -106.0212
23.923 | -106.0234
23.9267 | -106.0244
23.9305 | -106.0252
23.9342 | -106.0258
23.938 | -106.0245
23.9417 | -106.0224
23.9455 | -106.0214
23.9493 | -106.0199
23.953 | -106.018
23.9568 | -106.0153
23.9605 | -106.0111
23.9643 | -106.0078
23.9681 | -106.0035
23.9718 | -105.9995
23.9756 | -105.9969
23.9793 | -105.9914
23.9831 | -105.9839
23.9868 | -105.9751
23.9906 | -105.9662
23.9944 | -105.9575
23.9981 | -105.9481
24.0019 | -105.9404
24.0056 | -105.9315
24.0094 | -105.9214
24.0132 | -105.9131
24.0169 | -105.9029
24.0207 | -105.8932
24.0244 | -105.8836
24.0282 | -105.8752
24.0319 | -105.8672
24.0357 | -105.86
24.0395 | -105.851
24.0432 | -105.8419
24.047 | -105.8317
24.0507 | -105.8229
24.0545 | -105.8157
24.0583 | -105.8097
24.062 | -105.8021
24.0658 | -105.7946
24.0695 | -105.7905
24.0733 | -105.7863
24.077 | -105.7807
24.0808 | -105.7753
24.0846 | -105.7715
24.0883 | -105.7675
24.0921 | -105.7631
24.0958 | -105.7594
24.0996 | -105.7551
24.1034 | -105.7512
24.1071 | -105.7477
24.1109 | -105.7439
24.1146 | -105.7408
24.1184 | -105.7404
24.1222 | -105.7413
24.1259 | -105.7422
24.1297 | -105.7396
24.1334 | -105.7357
24.1372 | -105.7344
24.1409 | -105.7335
24.1447 | -105.7332
24.1485 | -105.7327
24.1522 | -105.7315
24.156 | -105.7314
24.1597 | -105.73
24.1635 | -105.7297
24.1673 | -105.7293
24.171 | -105.7261
24.1748 | -105.7234
24.1785 | -105.7203
24.1823 | -105.7165
24.186 | -105.7132
24.1898 | -105.7104
24.1936 | -105.7067
24.1973 | -105.7032
24.2011 | -105.6978
24.2048 | -105.6945
24.2086 | -105.6907
24.2124 | -105.6847
24.2161 | -105.6786
24.2199 | -105.675
24.2236 | -105.6714
24.2274 | -105.6671
24.2311 | -105.6632
24.2349 | -105.6576
24.2387 | -105.6516
24.2424 | -105.645
24.2462 | -105.6409
24.2499 | -105.6378
24.2537 | -105.6313
24.2575 | -105.624
24.2612 | -105.6148
24.265 | -105.5937
24.2687 | -105.5781
24.2725 | -105.5553
24.2763 | -105.5196
24.28 | -105.4661
24.2838 | -105.4012
24.2875 | -105.3499
24.2913 | -105.308
24.295 | -105.2683
24.2988 | -105.2256
24.3026 | -105.1815
24.3063 | -105.1357
24.3101 | -105.0874
24.3138 | -105.0353
24.3176 | -104.981
24.3214 | -104.9248
24.3251 | -104.8661
24.3289 | -104.8041
24.3326 | -104.7412
24.3364 | -104.6763
24.3401 | -104.61
24.3439 | -104.5423
24.3477 | -104.4716
24.3514 | -104.3994
24.3552 | -104.3249
24.3589 | -104.2513
24.3627 | -104.1774
24.3665 | -104.104
24.3702 | -104.0304
24.374 | -103.957
24.3777 | -103.8851
24.3815 | -103.8149
24.3852 | -103.7464
24.389 | -103.6796
24.3928 | -103.6138
24.3965 | -103.5509
24.4003 | -103.5105
24.404 | -103.4571
24.4078 | -103.4067
24.4116 | -103.3597
24.4153 | -103.3155
24.4191 | -103.2747
24.4228 | -103.2394
24.4266 | -103.2093
24.4304 | -103.1832
24.4341 | -103.1614
24.4379 | -103.1445
24.4416 | -103.1308
24.4454 | -103.1225
24.4491 | -103.118
24.4529 | -103.1177
24.4567 | -103.1226
24.4604 | -103.1321
24.4642 | -103.1454
24.4679 | -103.1624
24.4717 | -103.1843
24.4755 | -103.2118
24.4792 | -103.2447
24.483 | -103.2805
24.4867 | -103.3188
24.4905 | -103.3587
24.4942 | -103.4036
24.498 | -103.4528
24.5018 | -103.5043
24.5055 | -103.5589
24.5093 | -103.6154
24.513 | -103.6732
24.5168 | -103.7338
24.5206 | -103.7963
24.5243 | -103.8597
24.5281 | -103.9239
24.5318 | -103.9887
24.5356 | -104.0531
24.5393 | -104.1196
24.5431 | -104.1864
24.5469 | -104.253
24.5506 | -104.317
24.5544 | -104.3793
24.5581 | -104.4412
24.5619 | -104.4984
24.5657 | -104.549
24.5694 | -104.5858
24.5732 | -104.6023
24.5769 | -104.5803
24.5807 | -104.53
24.5844 | -104.4476
24.5882 | -104.3715
24.592 | -104.294
24.5957 | -104.2121
24.5995 | -104.1472
24.6032 | -104.0868
24.607 | -104.0179
24.6108 | -103.9461
24.6145 | -103.8605
24.6183 | -103.7702
24.622 | -103.6793
24.6258 | -103.5865
24.6296 | -103.4881
24.6333 | -103.371
24.6371 | -103.2399
24.6408 | -103.1016
24.6446 | -102.9613
24.6483 | -102.8184
24.6521 | -102.6729
24.6559 | -102.5267
24.6596 | -102.3809
24.6634 | -102.2351
24.6671 | -102.0889
24.6709 | -101.9452
24.6747 | -101.804
24.6784 | -101.6641
24.6822 | -101.5275
24.6859 | -101.395
24.6897 | -101.266
24.6934 | -101.1414
24.6972 | -101.0224
24.701 | -100.9089
24.7047 | -100.8006
24.7085 | -100.6986
24.7122 | -100.6026
24.716 | -100.5129
24.7198 | -100.4295
24.7235 | -100.353
24.7273 | -100.2836
24.731 | -100.2214
24.7348 | -100.1672
24.7385 | -100.1206
24.7423 | -100.0808
24.7461 | -100.0482
24.7498 | -100.0235
24.7536 | -100.0065
24.7573 | -99.9973
24.7611 | -99.9954
24.7649 | -100.0013
24.7686 | -100.0154
24.7724 | -100.0375
24.7761 | -100.0665
24.7799 | -100.1031
24.7837 | -100.1465
24.7874 | -100.1974
24.7912 | -100.2564
24.7949 | -100.3223
24.7987 | -100.3954
24.8024 | -100.4752
24.8062 | -100.5608
24.81 | -100.6528
24.8137 | -100.7505
24.8175 | -100.8541
24.8212 | -100.9629
24.825 | -101.076
24.8288 | -101.1938
24.8325 | -101.3152
24.8363 | -101.4407
24.84 | -101.571
24.8438 | -101.7029
24.8475 | -101.8369
24.8513 | -101.9723
24.8551 | -102.1088
24.8588 | -102.2474
24.8626 | -102.387
24.8663 | -102.5263
24.8701 | -102.6643
24.8739 | -102.8016
24.8776 | -102.9371
24.8814 | -103.0695
24.8851 | -103.1983
24.8889 | -103.3235
24.8926 | -103.4453
24.8964 | -103.5626
24.9002 | -103.6773
24.9039 | -103.7878
24.9077 | -103.8925
24.9114 | -103.9951
24.9152 | -104.0952
24.919 | -104.1987
24.9227 | -104.31
24.9265 | -104.44
24.9302 | -104.576
24.934 | -104.7212
24.9378 | -104.8335
24.9415 | -104.8458
24.9453 | -104.9315
24.949 | -105.0018
24.9528 | -105.0359
24.9565 | -105.0498
24.9603 | -105.0631
24.9641 | -105.0755
24.9678 | -105.088
24.9716 | -105.0995
24.9753 | -105.1084
24.9791 | -105.1207
24.9829 | -105.1309
24.9866 | -105.1388
24.9904 | -105.1466
24.9941 | -105.1545
24.9979 | -105.1609
25.0016 | -105.1667
25.0054 | -105.1737
25.0092 | -105.1808
25.0129 | -105.1869
25.0167 | -105.1929
25.0204 | -105.2005
25.0242 | -105.2068
25.028 | -105.2139
25.0317 | -105.2227
25.0355 | -105.2297
25.0392 | -105.2331
25.043 | -105.2381
25.0467 | -105.244
25.0505 | -105.2494
25.0543 | -105.2545
25.058 | -105.2586
25.0618 | -105.2627
25.0655 | -105.2659
25.0693 | -105.2679
25.0731 | -105.2724
25.0768 | -105.2764
25.0806 | -105.2787
25.0843 | -105.2803
25.0881 | -105.282
25.0918 | -105.2837
25.0956 | -105.2871
25.0994 | -105.2919
25.1031 | -105.296
25.1069 | -105.3008
25.1106 | -105.3057
25.1144 | -105.3118
25.1182 | -105.3165
25.1219 | -105.3196
25.1257 | -105.322
25.1294 | -105.3273
25.1332 | -105.33
25.137 | -105.3345
25.1407 | -105.3407
25.1445 | -105.344
25.1482 | -105.3469
25.152 | -105.3489
25.1557 | -105.3526
25.1595 | -105.3571
25.1633 | -105.3619
25.167 | -105.3666
25.1708 | -105.3721
25.1745 | -105.3763
25.1783 | -105.3811
25.1821 | -105.3861
25.1858 | -105.3917
25.1896 | -105.3978
25.1933 | -105.4007
25.1971 | -105.4052
25.2008 | -105.4096
25.2046 | -105.4137
25.2084 | -105.4182
25.2121 | -105.4206
25.2159 | -105.4242
25.2196 | -105.4271
25.2234 | -105.4295
25.2272 | -105.4304
25.2309 | -105.4332
25.2347 | -105.4363
25.2384 | -105.4397
25.2422 | -105.4402
25.2459 | -105.4396
25.2497 | -105.4389
25.2535 | -105.4365
25.2572 | -105.4339
25.261 | -105.4326
25.2647 | -105.4305
25.2685 | -105.4283
25.2723 | -105.4265
25.276 | -105.4222
25.2798 | -105.4181
25.2835 | -105.4134
25.2873 | -105.4075
25.291 | -105.4026
25.2948 | -105.3983
25.2986 | -105.3936
25.3023 | -105.3871
25.3061 | -105.3801
25.3098 | -105.3749
25.3136 | -105.3694
25.3174 | -105.3633
25.3211 | -105.3571
25.3249 | -105.3499
25.3286 | -105.3427
25.3324 | -105.3355
25.3362 | -105.325
25.3399 | -105.3155
25.3437 | -105.3075
25.3474 | -105.2988
25.3512 | -105.2889
25.3549 | -105.2785
25.3587 | -105.2692
25.3625 | -105.2605
25.3662 | -105.2522
25.37 | -105.2428
25.3737 | -105.2334
25.3775 | -105.2253
25.3813 | -105.2164
25.385 | -105.2069
25.3888 | -105.1961
25.3925 | -105.1853
25.3963 | -105.175
25.4 | -105.1658
25.4038 | -105.1589
25.4076 | -105.15
25.4113 | -105.1416
25.4151 | -105.1342
25.4188 | -105.126
25.4226 | -105.1189
25.4264 | -105.1134
25.4301 | -105.1054
25.4339 | -105.0989
25.4376 | -105.093
25.4414 | -105.0861
25.4451 | -105.0814
25.4489 | -105.0772
25.4527 | -105.0731
25.4564 | -105.0686
25.4602 | -105.064
25.4639 | -105.0591
25.4677 | -105.0552
25.4715 | -105.0517
25.4752 | -105.0484
25.479 | -105.0461
25.4827 | -105.0434
25.4865 | -105.0404
25.4903 | -105.0377
25.494 | -105.0374
25.4978 | -105.0381
25.5015 | -105.0389
25.5053 | -105.0401
25.509 | -105.0395
25.5128 | -105.0405
25.5166 | -105.0417
25.5203 | -105.043
25.5241 | -105.0439
25.5278 | -105.0454
25.5316 | -105.047
25.5354 | -105.0497
25.5391 | -105.0526
25.5429 | -105.0554
25.5466 | -105.0585
25.5504 | -105.0613
25.5541 | -105.065
25.5579 | -105.0679
25.5617 | -105.0719
25.5654 | -105.0778
25.5692 | -105.0837
25.5729 | -105.0899
25.5767 | -105.0945
25.5805 | -105.1
25.5842 | -105.1059
25.588 | -105.1118
25.5917 | -105.1169
25.5955 | -105.1227
25.5992 | -105.1289
25.603 | -105.1356
25.6068 | -105.1431
25.6105 | -105.149
25.6143 | -105.1537
25.618 | -105.1595
25.6218 | -105.1668
25.6256 | -105.1727
25.6293 | -105.1773
25.6331 | -105.1824
25.6368 | -105.19
25.6406 | -105.1974
25.6443 | -105.2055
25.6481 | -105.2127
25.6519 | -105.2199
25.6556 | -105.2272
25.6594 | -105.2342
25.6631 | -105.2423
25.6669 | -105.2512
25.6707 | -105.2578
25.6744 | -105.2662
25.6782 | -105.2755
25.6819 | -105.284
25.6857 | -105.2915
25.6895 | -105.2985
25.6932 | -105.3049
25.697 | -105.3142
25.7007 | -105.3251
25.7045 | -105.3342
25.7082 | -105.3432
25.712 | -105.3539
25.7158 | -105.3653
25.7195 | -105.3756
25.7233 | -105.3837
25.727 | -105.3905
25.7308 | -105.4
25.7346 | -105.409
25.7383 | -105.4169
25.7421 | -105.4258
25.7458 | -105.4344
25.7496 | -105.4435
25.7533 | -105.4551
25.7571 | -105.4669
25.7609 | -105.4758
25.7646 | -105.4826
25.7684 | -105.4905
25.7721 | -105.4984
25.7759 | -105.5052
25.7797 | -105.5123
25.7834 | -105.5197
25.7872 | -105.5286
25.7909 | -105.5389
25.7947 | -105.5463
25.7984 | -105.5527
25.8022 | -105.5611
25.806 | -105.5671
25.8097 | -105.5748
25.8135 | -105.5807
25.8172 | -105.5849
25.821 | -105.5893
25.8248 | -105.5948
25.8285 | -105.6008
25.8323 | -105.6049
25.836 | -105.6112
25.8398 | -105.6159
25.8436 | -105.622
25.8473 | -105.6268
25.8511 | -105.6317
25.8548 | -105.6351
25.8586 | -105.6371
25.8623 | -105.6372
25.8661 | -105.6349
25.8699 | -105.635
25.8736 | -105.6386
25.8774 | -105.6412
25.8811 | -105.642
25.8849 | -105.6415
25.8887 | -105.6408
25.8924 | -105.6416
25.8962 | -105.642
25.8999 | -105.6422
25.9037 | -105.6416
25.9074 | -105.6389
25.9112 | -105.6357
25.915 | -105.6313
25.9187 | -105.6287
25.9225 | -105.626
25.9262 | -105.6227
25.93 | -105.618
25.9338 | -105.6124
25.9375 | -105.6082
25.9413 | -105.6052
25.945 | -105.6015
25.9488 | -105.5983
25.9525 | -105.5964
25.9563 | -105.5933
25.9601 | -105.5894
25.9638 | -105.5855
25.9676 | -105.5815
25.9713 | -105.5769
25.9751 | -105.5741
25.9789 | -105.5726
25.9826 | -105.5698
25.9864 | -105.5658
25.9901 | -105.5621
25.9939 | -105.5604
25.9977 | -105.5585
26.0014 | -105.5555
26.0052 | -105.5536
26.0089 | -105.552
26.0127 | -105.5507
26.0164 | -105.5508
26.0202 | -105.551
26.024 | -105.5571
26.0277 | -105.558
26.0315 | -105.5604
26.0352 | -105.5624
26.039 | -105.5659
26.0428 | -105.5697
26.0465 | -105.5743
26.0503 | -105.579
26.054 | -105.5844
26.0578 | -105.5903
26.0615 | -105.5979
26.0653 | -105.6051
26.0691 | -105.6115
26.0728 | -105.6185
26.0766 | -105.6265
26.0803 | -105.6359
26.0841 | -105.6463
26.0879 | -105.6571
26.0916 | -105.6686
26.0954 | -105.6813
26.0991 | -105.6934
26.1029 | -105.7057
26.1066 | -105.7178
26.1104 | -105.731
26.1142 | -105.7442
26.1179 | -105.758
26.1217 | -105.7706
26.1254 | -105.7847
26.1292 | -105.7994
26.133 | -105.8123
26.1367 | -105.8271
26.1405 | -105.8427
26.1442 | -105.8578
26.148 | -105.8736
26.1517 | -105.8902
26.1555 | -105.9043
26.1593 | -105.9182
26.163 | -105.9322
26.1668 | -105.9484
26.1705 | -105.9639
26.1743 | -105.977
26.1781 | -105.9854
26.1818 | -105.9935
26.1856 | -106.0027
26.1893 | -106.0111
26.1931 | -106.0183
26.1969 | -106.0254
26.2006 | -106.0343
26.2044 | -106.0404
26.2081 | -106.0478
26.2119 | -106.0542
26.2156 | -106.0591
26.2194 | -106.0627
26.2232 | -106.0659
26.2269 | -106.0699
26.2307 | -106.074
26.2344 | -106.078
26.2382 | -106.0797
26.242 | -106.0795
26.2457 | -106.076
26.2495 | -106.075
26.2532 | -106.0758
26.257 | -106.0746
26.2607 | -106.0717
26.2645 | -106.0676
26.2683 | -106.063
26.272 | -106.0577
26.2758 | -106.0539
26.2795 | -106.0504
26.2833 | -106.0481
26.2871 | -106.0432
26.2908 | -106.0374
26.2946 | -106.0309
26.2983 | -106.0255
26.3021 | -106.0208
26.3058 | -106.0152
26.3096 | -106.0075
26.3134 | -106.0012
26.3171 | -105.9949
26.3209 | -105.9876
26.3246 | -105.98
26.3284 | -105.9744
26.3322 | -105.9705
26.3359 | -105.9683
26.3397 | -105.966
26.3434 | -105.964
26.3472 | -105.9614
26.351 | -105.9606
26.3547 | -105.9593
26.3585 | -105.9592
26.3622 | -105.9595
26.366 | -105.9601
26.3697 | -105.9613
26.3735 | -105.9624
26.3773 | -105.9645
26.381 | -105.9677
26.3848 | -105.9709
26.3885 | -105.9752
26.3923 | -105.9794
26.3961 | -105.9855
26.3998 | -105.9922
26.4036 | -105.9968
26.4073 | -106.0022
26.4111 | -106.01
26.4148 | -106.0184
26.4186 | -106.026
26.4224 | -106.0341
26.4261 | -106.0419
26.4299 | -106.0504
26.4336 | -106.0578
26.4374 | -106.0654
26.4412 | -106.0739
26.4449 | -106.0829
26.4487 | -106.0922
26.4524 | -106.1009
26.4562 | -106.1092
26.4599 | -106.1173
26.4637 | -106.1247
26.4675 | -106.1304
26.4712 | -106.1368
26.475 | -106.1441
26.4787 | -106.151
26.4825 | -106.1564
26.4863 | -106.1621
26.49 | -106.1678
26.4938 | -106.1733
26.4975 | -106.1773
26.5013 | -106.1828
26.5051 | -106.1899
26.5088 | -106.1949
26.5126 | -106.1986
26.5163 | -106.2028
26.5201 | -106.2067
26.5238 | -106.2096
26.5276 | -106.2131
26.5314 | -106.216
26.5351 | -106.2186
26.5389 | -106.2198
26.5426 | -106.2219
26.5464 | -106.2258
26.5502 | -106.2337
26.5539 | -106.2414
26.5577 | -106.2476
26.5614 | -106.253
26.5652 | -106.2566
26.5689 | -106.2587
26.5727 | -106.2612
26.5765 | -106.2628
26.5802 | -106.2646
26.584 | -106.2668
26.5877 | -106.2688
26.5915 | -106.2713
26.5953 | -106.2731
26.599 | -106.2745
26.6028 | -106.2773
26.6065 | -106.279
26.6103 | -106.2815
26.614 | -106.2858
26.6178 | -106.2902
26.6216 | -106.2968
26.6253 | -106.3016
26.6291 | -106.3034
26.6328 | -106.3087
26.6366 | -106.3141
26.6404 | -106.3191
26.6441 | -106.3245
26.6479 | -106.332
26.6516 | -106.3387
26.6554 | -106.3446
26.6592 | -106.3486
26.6629 | -106.3547
26.6667 | -106.3595
26.6704 | -106.365
26.6742 | -106.3698
26.6779 | -106.3753
26.6817 | -106.382
26.6855 | -106.39
26.6892 | -106.3966
26.693 | -106.4024
26.6967 | -106.4099
26.7005 | -106.4176
26.7043 | -106.4233
26.708 | -106.4283
26.7118 | -106.4333
26.7155 | -106.4384
26.7193 | -106.4427
26.723 | -106.449
26.7268 | -106.4545
26.7306 | -106.4595
26.7343 | -106.4642
26.7381 | -106.4679
26.7418 | -106.4707
26.7456 | -106.4717
26.7494 | -106.4732
26.7531 | -106.4744
26.7569 | -106.4766
26.7606 | -106.4792
26.7644 | -106.4816
26.7681 | -106.4845
26.7719 | -106.4849
26.7757 | -106.485
26.7794 | -106.4872
26.7832 | -106.4879
26.7869 | -106.4862
26.7907 | -106.4845
26.7945 | -106.4833
26.7982 | -106.4823
26.802 | -106.4808
26.8057 | -106.4778
26.8095 | -106.4753
26.8132 | -106.4724
26.817 | -106.4687
26.8208 | -106.464
26.8245 | -106.4616
26.8283 | -106.4584
26.832 | -106.4543
26.8358 | -106.4518
26.8396 | -106.45
26.8433 | -106.4493
26.8471 | -106.4477
26.8508 | -106.4461
26.8546 | -106.4437
26.8584 | -106.4416
26.8621 | -106.4395
26.8659 | -106.4357
26.8696 | -106.4322
26.8734 | -106.4313
26.8771 | -106.43
26.8809 | -106.4294
26.8847 | -106.4301
26.8884 | -106.4317
26.8922 | -106.4333
26.8959 | -106.4357
26.8997 | -106.4369
26.9035 | -106.438
26.9072 | -106.4407
26.911 | -106.4446
26.9147 | -106.4506
26.9185 | -106.4548
26.9222 | -106.457
26.926 | -106.4596
26.9298 | -106.4631
26.9335 | -106.4679
26.9373 | -106.4742
26.941 | -106.4811
26.9448 | -106.4883
26.9486 | -106.4957
26.9523 | -106.5029
26.9561 | -106.5107
26.9598 | -106.5174
26.9636 | -106.5235
26.9673 | -106.5312
26.9711 | -106.5397
26.9749 | -106.5482
26.9786 | -106.5571
26.9824 | -106.5662
26.9861 | -106.5764
26.9899 | -106.5866
26.9937 | -106.5944
26.9974 | -106.6019
27.0012 | -106.6103
27.0049 | -106.6199
27.0087 | -106.6314
27.0125 | -106.638
27.0162 | -106.6449
27.02 | -106.653
27.0237 | -106.6604
27.0275 | -106.6669
27.0312 | -106.6733
27.035 | -106.6797
27.0388 | -106.6877
27.0425 | -106.6951
27.0463 | -106.7019
27.05 | -106.7069
27.0538 | -106.7118
27.0576 | -106.717
27.0613 | -106.7193
27.0651 | -106.723
27.0688 | -106.7257
27.0726 | -106.7265
27.0763 | -106.726
27.0801 | -106.727
27.0839 | -106.7266
27.0876 | -106.7248
27.0914 | -106.72
27.0951 | -106.7169
27.0989 | -106.7155
27.1027 | -106.7116
27.1064 | -106.7068
27.1102 | -106.6989
27.1139 | -106.6946
27.1177 | -106.6881
27.1214 | -106.6814
27.1252 | -106.6755
27.129 | -106.671
27.1327 | -106.6652
27.1365 | -106.6573
27.1402 | -106.6484
27.144 | -106.6386
27.1478 | -106.6278
27.1515 | -106.6156
27.1553 | -106.606
27.159 | -106.5968
27.1628 | -106.5856
27.1665 | -106.5762
27.1703 | -106.5663
27.1741 | -106.5564
27.1778 | -106.5466
27.1816 | -106.5367
27.1853 | -106.5278
27.1891 | -106.5199
27.1929 | -106.51
27.1966 | -106.4996
27.2004 | -106.49
27.2041 | -106.4817
27.2079 | -106.4705
27.2117 | -106.4602
27.2154 | -106.4515
27.2192 | -106.4429
27.2229 | -106.4345
27.2267 | -106.4258
27.2304 | -106.4179
27.2342 | -106.4102
27.238 | -106.4031
27.2417 | -106.3958
27.2455 | -106.39
27.2492 | -106.3856
27.253 | -106.3802
27.2568 | -106.3736
27.2605 | -106.3675
27.2643 | -106.3619
27.268 | -106.3558
27.2718 | -106.3504
27.2755 | -106.3447
27.2793 | -106.3405
27.2831 | -106.3381
27.2868 | -106.3358
27.2906 | -106.333
27.2943 | -106.3308
27.2981 | -106.3286
27.3019 | -106.3271
27.3056 | -106.3254
27.3094 | -106.3236
27.3131 | -106.3229
27.3169 | -106.3203
27.3206 | -106.318
27.3244 | -106.3163
27.3282 | -106.3143
27.3319 | -106.3132
27.3357 | -106.3112
27.3394 | -106.3109
27.3432 | -106.3121
27.347 | -106.3139
27.3507 | -106.3139
27.3545 | -106.3145
27.3582 | -106.3164
27.362 | -106.3166
27.3658 | -106.3147
27.3695 | -106.3136
27.3733 | -106.3131
27.377 | -106.3137
27.3808 | -106.3141
27.3845 | -106.3142
27.3883 | -106.3144
27.3921 | -106.3161
27.3958 | -106.319
27.3996 | -106.321
27.4033 | -106.322
27.4071 | -106.3229
27.4109 | -106.3227
27.4146 | -106.3218
27.4184 | -106.3221
27.4221 | -106.3226
27.4259 | -106.3229
27.4296 | -106.3237
27.4334 | -106.325
27.4372 | -106.3246
27.4409 | -106.322
27.4447 | -106.3219
27.4484 | -106.322
27.4522 | -106.3223
27.456 | -106.3237
27.4597 | -106.3249
27.4635 | -106.3258
27.4672 | -106.3263
27.471 | -106.3269
27.4747 | -106.3287
27.4785 | -106.328
27.4823 | -106.3277
27.486 | -106.3278
27.4898 | -106.3275
27.4935 | -106.3259
27.4973 | -106.3245
27.5011 | -106.3244
27.5048 | -106.325
27.5086 | -106.3252
27.5123 | -106.3256
27.5161 | -106.3251
27.5199 | -106.3237
27.5236 | -106.3234
27.5274 | -106.323
27.5311 | -106.3223
27.5349 | -106.3222
27.5386 | -106.3227
27.5424 | -106.323
27.5462 | -106.3216
27.5499 | -106.3204
27.5537 | -106.3194
27.5574 | -106.3191
27.5612 | -106.3169
27.565 | -106.3151
27.5687 | -106.3143
27.5725 | -106.3117
27.5762 | -106.3084
27.58 | -106.3072
27.5837 | -106.3065
27.5875 | -106.3052
27.5913 | -106.3038
27.595 | -106.3035
27.5988 | -106.302
27.6025 | -106.2999
27.6063 | -106.2982
27.6101 | -106.2958
27.6138 | -106.2927
27.6176 | -106.2909
27.6213 | -106.2896
27.6251 | -106.2903
27.6288 | -106.2907
27.6326 | -106.2883
27.6364 | -106.285
27.6401 | -106.2837
27.6439 | -106.2831
27.6476 | -106.2809
27.6514 | -106.281
27.6552 | -106.2805
27.6589 | -106.2806
27.6627 | -106.2808
27.6664 | -106.2823
27.6702 | -106.2822
27.6739 | -106.282
27.6777 | -106.2828
27.6815 | -106.2823
27.6852 | -106.2817
27.689 | -106.2795
27.6927 | -106.2804
27.6965 | -106.2815
27.7003 | -106.281
27.704 | -106.2815
27.7078 | -106.2855
27.7115 | -106.2881
27.7153 | -106.2907
27.7191 | -106.2936
27.7228 | -106.295
27.7266 | -106.2973
27.7303 | -106.2998
27.7341 | -106.3024
27.7378 | -106.3044
27.7416 | -106.3062
27.7454 | -106.3082
27.7491 | -106.3106
27.7529 | -106.3151
27.7566 | -106.3206
27.7604 | -106.3259
27.7642 | -106.3294
27.7679 | -106.3338
27.7717 | -106.3388
27.7754 | -106.3429
27.7792 | -106.3448
27.7829 | -106.3463
27.7867 | -106.3478
27.7905 | -106.3506
27.7942 | -106.3555
27.798 | -106.3595
27.8017 | -106.3624
27.8055 | -106.3655
27.8093 | -106.3677
27.813 | -106.3688
27.8168 | -106.3701
27.8205 | -106.3724
27.8243 | -106.3745
27.828 | -106.3768
27.8318 | -106.3801
27.8356 | -106.3823
27.8393 | -106.3824
27.8431 | -106.3827
27.8468 | -106.3818
27.8506 | -106.3808
27.8544 | -106.3784
27.8581 | -106.3754
27.8619 | -106.3697
27.8656 | -106.3649
27.8694 | -106.3605
27.8731 | -106.3558
27.8769 | -106.3522
27.8807 | -106.3491
27.8844 | -106.3455
27.8882 | -106.3424
27.8919 | -106.337
27.8957 | -106.3306
27.8995 | -106.3231
27.9032 | -106.3165
27.907 | -106.3104
27.9107 | -106.305
27.9145 | -106.2998
27.9183 | -106.2916
27.922 | -106.2805
27.9258 | -106.2599
27.9295 | -106.2318
27.9333 | -106.2048
27.937 | -106.1873
27.9408 | -106.1736
27.9446 | -106.1592
27.9483 | -106.1444
27.9521 | -106.1289
27.9558 | -106.1122
27.9596 | -106.0953
27.9634 | -106.077
27.9671 | -106.0584
27.9709 | -106.0392
27.9746 | -106.0198
27.9784 | -105.9986
27.9821 | -105.9756
27.9859 | -105.9522
27.9897 | -105.9306
27.9934 | -105.9068
27.9972 | -105.88
28.0009 | -105.8537