## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 274 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 279
Hood | 02-92B/P | 230
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1012.8 | hPa |  | PRC PRO 01
Humidity | 62.4 | %/Hr
Internal temperature of calibration bench | 23.95 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1014 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.8 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.92 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3060 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 21/08/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 24/08/2018
Technical Manager | Anthony HUE | Approval |  | 30/08/2018
CEO | Robert GLEMOT | Approval |  | 30/08/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5056 | 57.6008
0.012209 | -36.0735 | 38.4311
0.018313 | -35.0149 | 28.2465
0.024417 | -34.5466 | 22.0117
0.030521 | -34.4346 | 18.2337
0.036626 | -34.2887 | 15.1754
0.04273 | -34.1908 | 12.8143
0.048834 | -34.1429 | 11.2992
0.054938 | -34.0805 | 10.1667
0.061043 | -34.0224 | 9.0768
0.067147 | -34.0103 | 8.2578
0.073251 | -34.0093 | 7.4979
0.079355 | -34.0264 | 6.8297
0.08546 | -34.0409 | 6.2269
0.091564 | -34.0368 | 5.7808
0.097668 | -34.0184 | 5.5263
0.10377 | -33.9964 | 5.2784
0.10988 | -33.9793 | 5.0063
0.11598 | -33.97 | 4.7249
0.12209 | -33.9712 | 4.4523
0.12819 | -33.9798 | 4.1747
0.13429 | -33.9895 | 3.8864
0.1404 | -33.993 | 3.6051
0.1465 | -33.995 | 3.3464
0.15261 | -33.9926 | 3.1262
0.15871 | -33.9863 | 2.9654
0.16482 | -33.9783 | 2.8645
0.17092 | -33.9761 | 2.8261
0.17702 | -33.974 | 2.7549
0.18313 | -33.9731 | 2.677
0.18923 | -33.9724 | 2.5689
0.19534 | -33.9709 | 2.4428
0.20144 | -33.9695 | 2.3021
0.20754 | -33.967 | 2.1822
0.21365 | -33.9662 | 2.0684
0.21975 | -33.966 | 1.971
0.22586 | -33.9672 | 1.884
0.23196 | -33.9697 | 1.812
0.23807 | -33.9736 | 1.7503
0.24417 | -33.9772 | 1.6932
0.25027 | -33.9808 | 1.6323
0.25638 | -33.9828 | 1.5716
0.26248 | -33.9842 | 1.506
0.26859 | -33.9848 | 1.4411
0.27469 | -33.9849 | 1.3701
0.2808 | -33.9857 | 1.2957
0.2869 | -33.9862 | 1.2243
0.293 | -33.9869 | 1.1567
0.29911 | -33.9876 | 1.0959
0.30521 | -33.9884 | 1.0365
0.31132 | -33.9891 | 0.99834
0.31742 | -33.9903 | 0.95604
0.32353 | -33.9907 | 0.92189
0.32963 | -33.9917 | 0.89403
0.33573 | -33.9929 | 0.87516
0.34184 | -33.9945 | 0.85543
0.34794 | -33.9968 | 0.83343
0.35405 | -33.9991 | 0.81106
0.36015 | -34.0017 | 0.78717
0.36626 | -34.0038 | 0.76076
0.37236 | -34.0058 | 0.72485
0.37846 | -34.0079 | 0.68123
0.38457 | -34.0086 | 0.62367
0.39067 | -34.0094 | 0.57689
0.39678 | -34.0099 | 0.52471
0.40288 | -34.0101 | 0.47016
0.40899 | -34.0098 | 0.41767
0.41509 | -34.0089 | 0.36607
0.42119 | -34.0086 | 0.32263
0.4273 | -34.0084 | 0.27614
0.4334 | -34.0081 | 0.22978
0.43951 | -34.0082 | 0.18657
0.44561 | -34.0089 | 0.1468
0.45172 | -34.0104 | 0.11558
0.45782 | -34.011 | 0.078749
0.46392 | -34.0121 | 0.042936
0.47003 | -34.0131 | 0.010823
0.47613 | -34.0138 | -0.022812
0.48224 | -34.0143 | -0.056913
0.48834 | -34.0133 | -0.090904
0.49445 | -34.0131 | -0.12544
0.50055 | -34.0125 | -0.15692
0.50665 | -34.0118 | -0.1904
0.51276 | -34.0108 | -0.22379
0.51886 | -34.0098 | -0.25723
0.52497 | -34.0086 | -0.28099
0.53107 | -34.0079 | -0.30981
0.53717 | -34.0072 | -0.33741
0.54328 | -34.0065 | -0.35981
0.54938 | -34.0063 | -0.38198
0.55549 | -34.0063 | -0.40259
0.56159 | -34.0074 | -0.41435
0.5677 | -34.0077 | -0.4289
0.5738 | -34.0087 | -0.4452
0.5799 | -34.0094 | -0.46009
0.58601 | -34.0104 | -0.47034
0.59211 | -34.0119 | -0.47715
0.59822 | -34.0132 | -0.48697
0.60432 | -34.0146 | -0.49932
0.61043 | -34.0157 | -0.5077
0.61653 | -34.0168 | -0.51825
0.62263 | -34.0183 | -0.52789
0.62874 | -34.0188 | -0.53868
0.63484 | -34.0196 | -0.54814
0.64095 | -34.0199 | -0.55846
0.64705 | -34.0202 | -0.56829
0.65316 | -34.0204 | -0.58081
0.65926 | -34.0204 | -0.59645
0.66536 | -34.0193 | -0.61338
0.67147 | -34.0188 | -0.6263
0.67757 | -34.0184 | -0.64027
0.68368 | -34.0177 | -0.65871
0.68978 | -34.0169 | -0.67541
0.69589 | -34.0159 | -0.69393
0.70199 | -34.0149 | -0.71605
0.70809 | -34.0141 | -0.73465
0.7142 | -34.0133 | -0.75454
0.7203 | -34.0122 | -0.77314
0.72641 | -34.0115 | -0.79503
0.73251 | -34.0109 | -0.81455
0.73862 | -34.011 | -0.83598
0.74472 | -34.0107 | -0.85564
0.75082 | -34.0104 | -0.87732
0.75693 | -34.0107 | -0.8988
0.76303 | -34.0112 | -0.92046
0.76914 | -34.0129 | -0.94081
0.77524 | -34.0138 | -0.96022
0.78135 | -34.015 | -0.98004
0.78745 | -34.0163 | -0.9998
0.79355 | -34.0174 | -1.0195
0.79966 | -34.0191 | -1.0388
0.80576 | -34.0213 | -1.0584
0.81187 | -34.0228 | -1.0797
0.81797 | -34.0247 | -1.1007
0.82408 | -34.0263 | -1.1212
0.83018 | -34.028 | -1.1415
0.83628 | -34.0298 | -1.16
0.84239 | -34.0309 | -1.1803
0.84849 | -34.0323 | -1.2018
0.8546 | -34.0337 | -1.2214
0.8607 | -34.0349 | -1.242
0.86681 | -34.0359 | -1.2608
0.87291 | -34.0368 | -1.2794
0.87901 | -34.0372 | -1.2972
0.88512 | -34.0379 | -1.3152
0.89122 | -34.0386 | -1.3333
0.89733 | -34.039 | -1.3488
0.90343 | -34.0394 | -1.3662
0.90953 | -34.0395 | -1.3803
0.91564 | -34.0394 | -1.396
0.92174 | -34.0396 | -1.4108
0.92785 | -34.0396 | -1.4252
0.93395 | -34.0397 | -1.4388
0.94006 | -34.0398 | -1.4515
0.94616 | -34.0394 | -1.4631
0.95226 | -34.0395 | -1.4756
0.95837 | -34.0395 | -1.4892
0.96447 | -34.0397 | -1.5018
0.97058 | -34.0395 | -1.515
0.97668 | -34.0392 | -1.5281
0.98279 | -34.0389 | -1.5405
0.98889 | -34.0387 | -1.551
0.99499 | -34.0386 | -1.5642
1.0011 | -34.0384 | -1.5746
1.0072 | -34.038 | -1.5861
1.0133 | -34.0376 | -1.5967
1.0194 | -34.0372 | -1.6095
1.0255 | -34.0368 | -1.623
1.0316 | -34.0363 | -1.6347
1.0377 | -34.0359 | -1.6483
1.0438 | -34.0353 | -1.6608
1.0499 | -34.0345 | -1.6759
1.056 | -34.0338 | -1.69
1.0621 | -34.0332 | -1.7041
1.0682 | -34.0328 | -1.7165
1.0743 | -34.0321 | -1.7317
1.0805 | -34.0316 | -1.746
1.0866 | -34.0307 | -1.761
1.0927 | -34.0299 | -1.7754
1.0988 | -34.0293 | -1.7902
1.1049 | -34.0285 | -1.8043
1.111 | -34.0277 | -1.8196
1.1171 | -34.027 | -1.8342
1.1232 | -34.0266 | -1.8493
1.1293 | -34.0262 | -1.8634
1.1354 | -34.0259 | -1.8783
1.1415 | -34.0255 | -1.8937
1.1476 | -34.0249 | -1.9092
1.1537 | -34.0244 | -1.9238
1.1598 | -34.0245 | -1.9385
1.1659 | -34.0242 | -1.9518
1.172 | -34.024 | -1.9657
1.1781 | -34.0237 | -1.9796
1.1842 | -34.0236 | -1.9937
1.1903 | -34.0239 | -2.0046
1.1964 | -34.0238 | -2.0179
1.2025 | -34.0239 | -2.0318
1.2086 | -34.0242 | -2.0442
1.2147 | -34.0242 | -2.0566
1.2209 | -34.0245 | -2.0677
1.227 | -34.0252 | -2.0783
1.2331 | -34.0255 | -2.0901
1.2392 | -34.0258 | -2.1018
1.2453 | -34.0262 | -2.1127
1.2514 | -34.0266 | -2.1225
1.2575 | -34.0271 | -2.1332
1.2636 | -34.0279 | -2.1418
1.2697 | -34.0285 | -2.1508
1.2758 | -34.0291 | -2.1616
1.2819 | -34.0297 | -2.1713
1.288 | -34.0303 | -2.1818
1.2941 | -34.0311 | -2.19
1.3002 | -34.0318 | -2.1978
1.3063 | -34.0322 | -2.2066
1.3124 | -34.0329 | -2.2154
1.3185 | -34.0335 | -2.2241
1.3246 | -34.0342 | -2.2329
1.3307 | -34.035 | -2.2412
1.3368 | -34.0357 | -2.2495
1.3429 | -34.0363 | -2.2581
1.349 | -34.0368 | -2.2671
1.3551 | -34.0374 | -2.2761
1.3613 | -34.0379 | -2.284
1.3674 | -34.0385 | -2.2938
1.3735 | -34.0389 | -2.3032
1.3796 | -34.0394 | -2.312
1.3857 | -34.0397 | -2.3209
1.3918 | -34.0402 | -2.3308
1.3979 | -34.0407 | -2.3397
1.404 | -34.0412 | -2.3508
1.4101 | -34.0416 | -2.3609
1.4162 | -34.0419 | -2.3712
1.4223 | -34.0424 | -2.3823
1.4284 | -34.0428 | -2.3921
1.4345 | -34.0431 | -2.4034
1.4406 | -34.0434 | -2.4167
1.4467 | -34.0438 | -2.428
1.4528 | -34.0442 | -2.4392
1.4589 | -34.0447 | -2.4508
1.465 | -34.0449 | -2.4615
1.4711 | -34.0454 | -2.4727
1.4772 | -34.0459 | -2.4859
1.4833 | -34.0463 | -2.4986
1.4894 | -34.0468 | -2.5106
1.4955 | -34.0473 | -2.5224
1.5016 | -34.0479 | -2.5348
1.5078 | -34.0482 | -2.5485
1.5139 | -34.0486 | -2.5614
1.52 | -34.049 | -2.5739
1.5261 | -34.0493 | -2.5864
1.5322 | -34.0496 | -2.6002
1.5383 | -34.05 | -2.6126
1.5444 | -34.0505 | -2.6252
1.5505 | -34.051 | -2.6385
1.5566 | -34.0514 | -2.6514
1.5627 | -34.0519 | -2.665
1.5688 | -34.0525 | -2.6784
1.5749 | -34.0528 | -2.691
1.581 | -34.0531 | -2.7043
1.5871 | -34.0534 | -2.7179
1.5932 | -34.0538 | -2.7307
1.5993 | -34.0541 | -2.7448
1.6054 | -34.0545 | -2.7579
1.6115 | -34.0549 | -2.7719
1.6176 | -34.0551 | -2.7844
1.6237 | -34.0553 | -2.798
1.6298 | -34.0556 | -2.8105
1.6359 | -34.0559 | -2.8232
1.642 | -34.0563 | -2.8351
1.6482 | -34.0565 | -2.8481
1.6543 | -34.0568 | -2.8617
1.6604 | -34.0571 | -2.8736
1.6665 | -34.0571 | -2.8861
1.6726 | -34.0573 | -2.8986
1.6787 | -34.0576 | -2.9111
1.6848 | -34.0575 | -2.9236
1.6909 | -34.0578 | -2.9363
1.697 | -34.058 | -2.9485
1.7031 | -34.058 | -2.9611
1.7092 | -34.0579 | -2.9733
1.7153 | -34.0581 | -2.9852
1.7214 | -34.0579 | -2.9973
1.7275 | -34.0578 | -3.0098
1.7336 | -34.0577 | -3.0227
1.7397 | -34.0578 | -3.0349
1.7458 | -34.0578 | -3.0471
1.7519 | -34.0578 | -3.0592
1.758 | -34.0578 | -3.0711
1.7641 | -34.0579 | -3.0826
1.7702 | -34.0578 | -3.0955
1.7763 | -34.0578 | -3.107
1.7824 | -34.0577 | -3.1193
1.7885 | -34.0577 | -3.1316
1.7947 | -34.0578 | -3.1435
1.8008 | -34.0578 | -3.1548
1.8069 | -34.0578 | -3.1663
1.813 | -34.0578 | -3.1784
1.8191 | -34.0577 | -3.191
1.8252 | -34.0578 | -3.2036
1.8313 | -34.0579 | -3.2157
1.8374 | -34.058 | -3.2274
1.8435 | -34.0581 | -3.2392
1.8496 | -34.0581 | -3.252
1.8557 | -34.0583 | -3.2646
1.8618 | -34.0585 | -3.2766
1.8679 | -34.0586 | -3.2884
1.874 | -34.0587 | -3.3003
1.8801 | -34.0589 | -3.3124
1.8862 | -34.0591 | -3.3243
1.8923 | -34.0592 | -3.3362
1.8984 | -34.0593 | -3.3481
1.9045 | -34.0594 | -3.3595
1.9106 | -34.0594 | -3.3714
1.9167 | -34.0594 | -3.3834
1.9228 | -34.0596 | -3.3951
1.9289 | -34.0597 | -3.4067
1.9351 | -34.0598 | -3.4174
1.9412 | -34.06 | -3.4284
1.9473 | -34.0603 | -3.4403
1.9534 | -34.0605 | -3.4515
1.9595 | -34.0606 | -3.4631
1.9656 | -34.0609 | -3.475
1.9717 | -34.0611 | -3.4865
1.9778 | -34.0614 | -3.4977
1.9839 | -34.0616 | -3.5089
1.99 | -34.0618 | -3.5198
1.9961 | -34.0619 | -3.5317
2.0022 | -34.0621 | -3.5428
2.0083 | -34.0622 | -3.5538
2.0144 | -34.0624 | -3.5651
2.0205 | -34.0627 | -3.5766
2.0266 | -34.063 | -3.5877
2.0327 | -34.0632 | -3.599
2.0388 | -34.0636 | -3.6103
2.0449 | -34.0638 | -3.6215
2.051 | -34.064 | -3.6326
2.0571 | -34.0643 | -3.6438
2.0632 | -34.0645 | -3.655
2.0693 | -34.0647 | -3.6669
2.0754 | -34.0649 | -3.6781
2.0816 | -34.0652 | -3.69
2.0877 | -34.0654 | -3.7013
2.0938 | -34.0657 | -3.7131
2.0999 | -34.0659 | -3.7246
2.106 | -34.066 | -3.736
2.1121 | -34.0662 | -3.7478
2.1182 | -34.0664 | -3.7583
2.1243 | -34.0667 | -3.7689
2.1304 | -34.0668 | -3.7798
2.1365 | -34.0671 | -3.7913
2.1426 | -34.0673 | -3.8031
2.1487 | -34.0676 | -3.8141
2.1548 | -34.0679 | -3.8257
2.1609 | -34.0683 | -3.8369
2.167 | -34.0687 | -3.848
2.1731 | -34.069 | -3.8597
2.1792 | -34.0691 | -3.8706
2.1853 | -34.0695 | -3.8822
2.1914 | -34.0697 | -3.8934
2.1975 | -34.0701 | -3.9047
2.2036 | -34.0703 | -3.9165
2.2097 | -34.0705 | -3.9277
2.2158 | -34.0708 | -3.9381
2.222 | -34.071 | -3.9484
2.2281 | -34.0713 | -3.9596
2.2342 | -34.0715 | -3.9707
2.2403 | -34.0717 | -3.9812
2.2464 | -34.0719 | -3.9918
2.2525 | -34.0721 | -4.003
2.2586 | -34.0724 | -4.0134
2.2647 | -34.0727 | -4.0236
2.2708 | -34.073 | -4.0347
2.2769 | -34.0733 | -4.0451
2.283 | -34.0734 | -4.0561
2.2891 | -34.0736 | -4.0668
2.2952 | -34.0737 | -4.0771
2.3013 | -34.0739 | -4.0878
2.3074 | -34.0741 | -4.0981
2.3135 | -34.0743 | -4.1093
2.3196 | -34.0745 | -4.1193
2.3257 | -34.0747 | -4.1297
2.3318 | -34.0748 | -4.1405
2.3379 | -34.0749 | -4.1511
2.344 | -34.075 | -4.1623
2.3501 | -34.0752 | -4.1728
2.3562 | -34.0753 | -4.1831
2.3623 | -34.0754 | -4.1933
2.3685 | -34.0755 | -4.2037
2.3746 | -34.0756 | -4.2147
2.3807 | -34.0758 | -4.225
2.3868 | -34.076 | -4.2356
2.3929 | -34.0761 | -4.2461
2.399 | -34.0762 | -4.2564
2.4051 | -34.0764 | -4.2667
2.4112 | -34.0765 | -4.2772
2.4173 | -34.0767 | -4.2877
2.4234 | -34.0767 | -4.2983
2.4295 | -34.0769 | -4.309
2.4356 | -34.077 | -4.3196
2.4417 | -34.0771 | -4.3306
2.4478 | -34.0772 | -4.3413
2.4539 | -34.0774 | -4.3514
2.46 | -34.0774 | -4.3627
2.4661 | -34.0776 | -4.3736
2.4722 | -34.0777 | -4.3837
2.4783 | -34.0778 | -4.394
2.4844 | -34.0778 | -4.4048
2.4905 | -34.0779 | -4.4157
2.4966 | -34.0781 | -4.4262
2.5027 | -34.0781 | -4.4366
2.5089 | -34.0783 | -4.4473
2.515 | -34.0783 | -4.458
2.5211 | -34.0783 | -4.4686
2.5272 | -34.0784 | -4.4794
2.5333 | -34.0785 | -4.4902
2.5394 | -34.0787 | -4.5008
2.5455 | -34.0788 | -4.512
2.5516 | -34.0789 | -4.5222
2.5577 | -34.0791 | -4.5335
2.5638 | -34.0792 | -4.5451
2.5699 | -34.0793 | -4.5564
2.576 | -34.0793 | -4.5679
2.5821 | -34.0794 | -4.579
2.5882 | -34.0795 | -4.59
2.5943 | -34.0796 | -4.6007
2.6004 | -34.0798 | -4.6115
2.6065 | -34.0799 | -4.6218
2.6126 | -34.08 | -4.6322
2.6187 | -34.0801 | -4.6426
2.6248 | -34.0803 | -4.6533
2.6309 | -34.0804 | -4.6646
2.637 | -34.0806 | -4.6761
2.6431 | -34.0808 | -4.6865
2.6492 | -34.0809 | -4.6977
2.6554 | -34.0811 | -4.7085
2.6615 | -34.0812 | -4.7194
2.6676 | -34.0814 | -4.7307
2.6737 | -34.0817 | -4.7417
2.6798 | -34.0819 | -4.7527
2.6859 | -34.0821 | -4.7639
2.692 | -34.0823 | -4.7747
2.6981 | -34.0825 | -4.7857
2.7042 | -34.0827 | -4.7967
2.7103 | -34.0828 | -4.8072
2.7164 | -34.0831 | -4.8181
2.7225 | -34.0834 | -4.8293
2.7286 | -34.0836 | -4.8402
2.7347 | -34.0838 | -4.8507
2.7408 | -34.084 | -4.8607
2.7469 | -34.0842 | -4.8718
2.753 | -34.0845 | -4.8823
2.7591 | -34.0848 | -4.8935
2.7652 | -34.085 | -4.9049
2.7713 | -34.0852 | -4.9155
2.7774 | -34.0854 | -4.9263
2.7835 | -34.0856 | -4.9372
2.7896 | -34.0858 | -4.9477
2.7958 | -34.086 | -4.9583
2.8019 | -34.0863 | -4.9686
2.808 | -34.0866 | -4.9793
2.8141 | -34.0869 | -4.9904
2.8202 | -34.0872 | -5.0012
2.8263 | -34.0875 | -5.012
2.8324 | -34.0878 | -5.0232
2.8385 | -34.0882 | -5.034
2.8446 | -34.0885 | -5.0446
2.8507 | -34.0889 | -5.0553
2.8568 | -34.0892 | -5.0657
2.8629 | -34.0894 | -5.0761
2.869 | -34.0897 | -5.0871
2.8751 | -34.09 | -5.0977
2.8812 | -34.0904 | -5.1089
2.8873 | -34.0907 | -5.1196
2.8934 | -34.091 | -5.1305
2.8995 | -34.0913 | -5.1415
2.9056 | -34.0916 | -5.1525
2.9117 | -34.0919 | -5.1634
2.9178 | -34.0922 | -5.174
2.9239 | -34.0926 | -5.1854
2.93 | -34.093 | -5.1967
2.9361 | -34.0933 | -5.2073
2.9423 | -34.0936 | -5.2179
2.9484 | -34.094 | -5.2288
2.9545 | -34.0943 | -5.2396
2.9606 | -34.0947 | -5.2509
2.9667 | -34.095 | -5.2618
2.9728 | -34.0953 | -5.2729
2.9789 | -34.0958 | -5.2839
2.985 | -34.0961 | -5.295
2.9911 | -34.0965 | -5.3061
2.9972 | -34.0969 | -5.317
3.0033 | -34.0973 | -5.3283
3.0094 | -34.0976 | -5.3393
3.0155 | -34.098 | -5.35
3.0216 | -34.0984 | -5.3611
3.0277 | -34.0987 | -5.3726
3.0338 | -34.099 | -5.3838
3.0399 | -34.0994 | -5.3949
3.046 | -34.0997 | -5.4063
3.0521 | -34.1 | -5.4172
3.0582 | -34.1003 | -5.429
3.0643 | -34.1006 | -5.4402
3.0704 | -34.1009 | -5.4518
3.0765 | -34.1012 | -5.4627
3.0827 | -34.1014 | -5.4738
3.0888 | -34.1017 | -5.4848
3.0949 | -34.1021 | -5.4959
3.101 | -34.1024 | -5.5071
3.1071 | -34.1027 | -5.5187
3.1132 | -34.1031 | -5.5302
3.1193 | -34.1034 | -5.5415
3.1254 | -34.1037 | -5.5526
3.1315 | -34.1041 | -5.5643
3.1376 | -34.1043 | -5.5754
3.1437 | -34.1046 | -5.5868
3.1498 | -34.105 | -5.5986
3.1559 | -34.1052 | -5.6099
3.162 | -34.1055 | -5.6217
3.1681 | -34.1057 | -5.6334
3.1742 | -34.106 | -5.6447
3.1803 | -34.1063 | -5.6561
3.1864 | -34.1066 | -5.6671
3.1925 | -34.1069 | -5.6785
3.1986 | -34.1071 | -5.6901
3.2047 | -34.1074 | -5.7019
3.2108 | -34.1076 | -5.7128
3.2169 | -34.1079 | -5.7243
3.223 | -34.1081 | -5.7357
3.2292 | -34.1083 | -5.7473
3.2353 | -34.1085 | -5.7579
3.2414 | -34.1087 | -5.7691
3.2475 | -34.1089 | -5.7803
3.2536 | -34.1091 | -5.7915
3.2597 | -34.1094 | -5.8023
3.2658 | -34.1095 | -5.8131
3.2719 | -34.1097 | -5.8239
3.278 | -34.11 | -5.8353
3.2841 | -34.1102 | -5.8461
3.2902 | -34.1104 | -5.8573
3.2963 | -34.1106 | -5.8686
3.3024 | -34.1107 | -5.8795
3.3085 | -34.111 | -5.8908
3.3146 | -34.1111 | -5.9018
3.3207 | -34.1113 | -5.9131
3.3268 | -34.1115 | -5.9243
3.3329 | -34.1117 | -5.9349
3.339 | -34.1118 | -5.9455
3.3451 | -34.112 | -5.9565
3.3512 | -34.1122 | -5.967
3.3573 | -34.1124 | -5.978
3.3634 | -34.1125 | -5.989
3.3696 | -34.1127 | -6.0001
3.3757 | -34.1128 | -6.0107
3.3818 | -34.113 | -6.0214
3.3879 | -34.1132 | -6.0323
3.394 | -34.1133 | -6.0435
3.4001 | -34.1135 | -6.0542
3.4062 | -34.1136 | -6.0649
3.4123 | -34.1137 | -6.0753
3.4184 | -34.1139 | -6.0864
3.4245 | -34.1141 | -6.0971
3.4306 | -34.1142 | -6.1076
3.4367 | -34.1143 | -6.118
3.4428 | -34.1145 | -6.1282
3.4489 | -34.1146 | -6.139
3.455 | -34.1147 | -6.1497
3.4611 | -34.1149 | -6.1608
3.4672 | -34.115 | -6.1717
3.4733 | -34.1152 | -6.1826
3.4794 | -34.1154 | -6.1933
3.4855 | -34.1155 | -6.2036
3.4916 | -34.1157 | -6.2138
3.4977 | -34.1158 | -6.2249
3.5038 | -34.1159 | -6.2356
3.5099 | -34.116 | -6.2458
3.5161 | -34.1162 | -6.2561
3.5222 | -34.1163 | -6.2664
3.5283 | -34.1164 | -6.277
3.5344 | -34.1165 | -6.2877
3.5405 | -34.1167 | -6.2981
3.5466 | -34.1168 | -6.3089
3.5527 | -34.1169 | -6.3197
3.5588 | -34.117 | -6.3302
3.5649 | -34.1171 | -6.3408
3.571 | -34.1172 | -6.3513
3.5771 | -34.1173 | -6.3622
3.5832 | -34.1175 | -6.372
3.5893 | -34.1177 | -6.3825
3.5954 | -34.1178 | -6.3932
3.6015 | -34.118 | -6.404
3.6076 | -34.1182 | -6.4144
3.6137 | -34.1183 | -6.4244
3.6198 | -34.1186 | -6.4346
3.6259 | -34.1187 | -6.4454
3.632 | -34.1189 | -6.4559
3.6381 | -34.119 | -6.4666
3.6442 | -34.1192 | -6.477
3.6503 | -34.1194 | -6.4871
3.6565 | -34.1196 | -6.4978
3.6626 | -34.1197 | -6.5082
3.6687 | -34.1199 | -6.5188
3.6748 | -34.1202 | -6.5291
3.6809 | -34.1204 | -6.5397
3.687 | -34.1205 | -6.5505
3.6931 | -34.1208 | -6.5606
3.6992 | -34.121 | -6.5712
3.7053 | -34.1212 | -6.5817
3.7114 | -34.1213 | -6.5923
3.7175 | -34.1215 | -6.6029
3.7236 | -34.1218 | -6.6129
3.7297 | -34.122 | -6.6235
3.7358 | -34.1222 | -6.6343
3.7419 | -34.1224 | -6.645
3.748 | -34.1226 | -6.6555
3.7541 | -34.1228 | -6.666
3.7602 | -34.123 | -6.6764
3.7663 | -34.1232 | -6.6868
3.7724 | -34.1235 | -6.6974
3.7785 | -34.1237 | -6.7077
3.7846 | -34.1239 | -6.7181
3.7907 | -34.1242 | -6.7289
3.7969 | -34.1244 | -6.7396
3.803 | -34.1246 | -6.7497
3.8091 | -34.1248 | -6.7601
3.8152 | -34.1251 | -6.7706
3.8213 | -34.1253 | -6.7813
3.8274 | -34.1256 | -6.7916
3.8335 | -34.1259 | -6.8023
3.8396 | -34.1262 | -6.8128
3.8457 | -34.1265 | -6.8235
3.8518 | -34.1268 | -6.8339
3.8579 | -34.1271 | -6.8444
3.864 | -34.1273 | -6.855
3.8701 | -34.1276 | -6.8659
3.8762 | -34.1279 | -6.8763
3.8823 | -34.1281 | -6.887
3.8884 | -34.1284 | -6.8976
3.8945 | -34.1287 | -6.9083
3.9006 | -34.129 | -6.9192
3.9067 | -34.1292 | -6.9302
3.9128 | -34.1295 | -6.9409
3.9189 | -34.1298 | -6.9513
3.925 | -34.13 | -6.962
3.9311 | -34.1303 | -6.9721
3.9372 | -34.1305 | -6.9829
3.9434 | -34.1308 | -6.9935
3.9495 | -34.1312 | -7.0041
3.9556 | -34.1315 | -7.0146
3.9617 | -34.1317 | -7.0247
3.9678 | -34.1321 | -7.035
3.9739 | -34.1323 | -7.0456
3.98 | -34.1325 | -7.0562
3.9861 | -34.1328 | -7.0668
3.9922 | -34.133 | -7.0774
3.9983 | -34.1333 | -7.0876
4.0044 | -34.1335 | -7.0977
4.0105 | -34.1338 | -7.1083
4.0166 | -34.1341 | -7.119
4.0227 | -34.1344 | -7.1294
4.0288 | -34.1346 | -7.1404
4.0349 | -34.1348 | -7.1507
4.041 | -34.1351 | -7.1612
4.0471 | -34.1353 | -7.1719
4.0532 | -34.1355 | -7.1821
4.0593 | -34.1358 | -7.1928
4.0654 | -34.1361 | -7.2031
4.0715 | -34.1363 | -7.2141
4.0776 | -34.1366 | -7.2246
4.0838 | -34.1369 | -7.2355
4.0899 | -34.1372 | -7.2458
4.096 | -34.1374 | -7.2563
4.1021 | -34.1376 | -7.2666
4.1082 | -34.1379 | -7.2769
4.1143 | -34.1381 | -7.287
4.1204 | -34.1384 | -7.2973
4.1265 | -34.1386 | -7.3075
4.1326 | -34.1389 | -7.3177
4.1387 | -34.1392 | -7.3281
4.1448 | -34.1394 | -7.3387
4.1509 | -34.1396 | -7.3491
4.157 | -34.1398 | -7.36
4.1631 | -34.1401 | -7.3707
4.1692 | -34.1403 | -7.381
4.1753 | -34.1406 | -7.3914
4.1814 | -34.1408 | -7.4015
4.1875 | -34.141 | -7.4122
4.1936 | -34.1413 | -7.4228
4.1997 | -34.1415 | -7.4328
4.2058 | -34.1417 | -7.4436
4.2119 | -34.1419 | -7.4538
4.218 | -34.1421 | -7.464
4.2241 | -34.1423 | -7.4742
4.2303 | -34.1426 | -7.4844
4.2364 | -34.1428 | -7.4948
4.2425 | -34.143 | -7.5053
4.2486 | -34.1432 | -7.5158
4.2547 | -34.1433 | -7.5259
4.2608 | -34.1435 | -7.5364
4.2669 | -34.1437 | -7.5465
4.273 | -34.1439 | -7.557
4.2791 | -34.1441 | -7.5673
4.2852 | -34.1443 | -7.5779
4.2913 | -34.1445 | -7.5879
4.2974 | -34.1447 | -7.5986
4.3035 | -34.1449 | -7.6092
4.3096 | -34.1451 | -7.6197
4.3157 | -34.1453 | -7.6299
4.3218 | -34.1455 | -7.6401
4.3279 | -34.1457 | -7.6505
4.334 | -34.1458 | -7.6607
4.3401 | -34.146 | -7.6706
4.3462 | -34.1463 | -7.6811
4.3523 | -34.1465 | -7.691
4.3584 | -34.1466 | -7.7009
4.3645 | -34.1468 | -7.7111
4.3707 | -34.147 | -7.7211
4.3768 | -34.1472 | -7.7314
4.3829 | -34.1474 | -7.7419
4.389 | -34.1476 | -7.7522
4.3951 | -34.1478 | -7.7622
4.4012 | -34.148 | -7.7728
4.4073 | -34.1481 | -7.7828
4.4134 | -34.1483 | -7.793
4.4195 | -34.1486 | -7.8032
4.4256 | -34.1488 | -7.8133
4.4317 | -34.149 | -7.8234
4.4378 | -34.1492 | -7.8333
4.4439 | -34.1493 | -7.8432
4.45 | -34.1495 | -7.8534
4.4561 | -34.1496 | -7.8635
4.4622 | -34.1498 | -7.8737
4.4683 | -34.15 | -7.8839
4.4744 | -34.1502 | -7.8944
4.4805 | -34.1504 | -7.9042
4.4866 | -34.1506 | -7.9144
4.4927 | -34.1508 | -7.9247
4.4988 | -34.151 | -7.9348
4.5049 | -34.1512 | -7.9446
4.511 | -34.1514 | -7.9545
4.5172 | -34.1516 | -7.9646
4.5233 | -34.1518 | -7.9751
4.5294 | -34.152 | -7.9851
4.5355 | -34.1522 | -7.9953
4.5416 | -34.1524 | -8.0054
4.5477 | -34.1526 | -8.0155
4.5538 | -34.1528 | -8.0253
4.5599 | -34.153 | -8.0353
4.566 | -34.1532 | -8.0455
4.5721 | -34.1534 | -8.0556
4.5782 | -34.1537 | -8.066
4.5843 | -34.1539 | -8.0761
4.5904 | -34.1541 | -8.0863
4.5965 | -34.1543 | -8.0963
4.6026 | -34.1545 | -8.1068
4.6087 | -34.1547 | -8.1168
4.6148 | -34.1549 | -8.1273
4.6209 | -34.1551 | -8.1373
4.627 | -34.1554 | -8.1476
4.6331 | -34.1555 | -8.1577
4.6392 | -34.1557 | -8.1678
4.6453 | -34.156 | -8.178
4.6514 | -34.1562 | -8.1881
4.6576 | -34.1564 | -8.1986
4.6637 | -34.1565 | -8.2085
4.6698 | -34.1567 | -8.2184
4.6759 | -34.1569 | -8.2289
4.682 | -34.1571 | -8.239
4.6881 | -34.1573 | -8.2491
4.6942 | -34.1575 | -8.2592
4.7003 | -34.1577 | -8.2694
4.7064 | -34.1579 | -8.2794
4.7125 | -34.1582 | -8.2899
4.7186 | -34.1583 | -8.2998
4.7247 | -34.1586 | -8.31
4.7308 | -34.1588 | -8.3201
4.7369 | -34.159 | -8.3304
4.743 | -34.1593 | -8.3409
4.7491 | -34.1595 | -8.3514
4.7552 | -34.1597 | -8.3617
4.7613 | -34.1599 | -8.372
4.7674 | -34.1601 | -8.382
4.7735 | -34.1602 | -8.3926
4.7796 | -34.1605 | -8.403
4.7857 | -34.1608 | -8.4134
4.7918 | -34.161 | -8.4233
4.7979 | -34.1612 | -8.4334
4.8041 | -34.1615 | -8.4437
4.8102 | -34.1617 | -8.4536
4.8163 | -34.1619 | -8.4637
4.8224 | -34.1621 | -8.4739
4.8285 | -34.1624 | -8.4838
4.8346 | -34.1626 | -8.4938
4.8407 | -34.1628 | -8.5042
4.8468 | -34.1631 | -8.5146
4.8529 | -34.1633 | -8.5251
4.859 | -34.1636 | -8.5356
4.8651 | -34.1638 | -8.5457
4.8712 | -34.164 | -8.5557
4.8773 | -34.1642 | -8.566
4.8834 | -34.1644 | -8.5762
4.8895 | -34.1647 | -8.5862
4.8956 | -34.1649 | -8.5965
4.9017 | -34.165 | -8.6066
4.9078 | -34.1653 | -8.6171
4.9139 | -34.1655 | -8.627
4.92 | -34.1658 | -8.6369
4.9261 | -34.1661 | -8.6468
4.9322 | -34.1663 | -8.657
4.9383 | -34.1666 | -8.6672
4.9445 | -34.1668 | -8.6774
4.9506 | -34.167 | -8.6874
4.9567 | -34.1673 | -8.6975
4.9628 | -34.1675 | -8.7077
4.9689 | -34.1677 | -8.7182
4.975 | -34.168 | -8.7287
4.9811 | -34.1682 | -8.7386
4.9872 | -34.1685 | -8.7487
4.9933 | -34.1688 | -8.7592
4.9994 | -34.1691 | -8.7698
5.0055 | -34.1693 | -8.7801
5.0116 | -34.1696 | -8.7907
5.0177 | -34.1699 | -8.8012
5.0238 | -34.1702 | -8.8117
5.0299 | -34.1705 | -8.8223
5.036 | -34.1708 | -8.8325
5.0421 | -34.171 | -8.8427
5.0482 | -34.1713 | -8.8529
5.0543 | -34.1715 | -8.8629
5.0604 | -34.1718 | -8.8732
5.0665 | -34.1721 | -8.8832
5.0726 | -34.1724 | -8.8938
5.0787 | -34.1727 | -8.9042
5.0848 | -34.1729 | -8.9146
5.091 | -34.1732 | -8.9247
5.0971 | -34.1735 | -8.9346
5.1032 | -34.1738 | -8.9453
5.1093 | -34.174 | -8.9555
5.1154 | -34.1743 | -8.9656
5.1215 | -34.1745 | -8.9757
5.1276 | -34.1748 | -8.9858
5.1337 | -34.1751 | -8.9958
5.1398 | -34.1754 | -9.0058
5.1459 | -34.1757 | -9.0164
5.152 | -34.1759 | -9.0268
5.1581 | -34.1762 | -9.0372
5.1642 | -34.1765 | -9.0475
5.1703 | -34.1768 | -9.0577
5.1764 | -34.1771 | -9.0678
5.1825 | -34.1773 | -9.0779
5.1886 | -34.1776 | -9.0878
5.1947 | -34.1779 | -9.098
5.2008 | -34.1782 | -9.1084
5.2069 | -34.1785 | -9.1188
5.213 | -34.1788 | -9.1288
5.2191 | -34.1791 | -9.1392
5.2252 | -34.1793 | -9.1495
5.2314 | -34.1796 | -9.16
5.2375 | -34.1798 | -9.1699
5.2436 | -34.1801 | -9.1802
5.2497 | -34.1804 | -9.1902
5.2558 | -34.1807 | -9.2008
5.2619 | -34.181 | -9.2113
5.268 | -34.1813 | -9.2218
5.2741 | -34.1816 | -9.2321
5.2802 | -34.1819 | -9.2427
5.2863 | -34.1822 | -9.2529
5.2924 | -34.1825 | -9.2633
5.2985 | -34.1828 | -9.2737
5.3046 | -34.1831 | -9.284
5.3107 | -34.1834 | -9.2944
5.3168 | -34.1837 | -9.3048
5.3229 | -34.184 | -9.3151
5.329 | -34.1843 | -9.3254
5.3351 | -34.1846 | -9.3354
5.3412 | -34.1848 | -9.3457
5.3473 | -34.1851 | -9.356
5.3534 | -34.1854 | -9.3664
5.3595 | -34.1857 | -9.3767
5.3656 | -34.186 | -9.3869
5.3717 | -34.1863 | -9.3975
5.3779 | -34.1866 | -9.4081
5.384 | -34.1869 | -9.4184
5.3901 | -34.1872 | -9.429
5.3962 | -34.1875 | -9.4391
5.4023 | -34.1878 | -9.4496
5.4084 | -34.1881 | -9.4598
5.4145 | -34.1884 | -9.4699
5.4206 | -34.1887 | -9.4802
5.4267 | -34.189 | -9.4908
5.4328 | -34.1893 | -9.5012
5.4389 | -34.1896 | -9.5112
5.445 | -34.1899 | -9.5218
5.4511 | -34.1902 | -9.5319
5.4572 | -34.1905 | -9.5423
5.4633 | -34.1908 | -9.5526
5.4694 | -34.191 | -9.563
5.4755 | -34.1913 | -9.573
5.4816 | -34.1916 | -9.5834
5.4877 | -34.192 | -9.5936
5.4938 | -34.1922 | -9.6036
5.4999 | -34.1925 | -9.6138
5.506 | -34.1929 | -9.624
5.5121 | -34.1931 | -9.6342
5.5183 | -34.1934 | -9.6442
5.5244 | -34.1937 | -9.6547
5.5305 | -34.194 | -9.6649
5.5366 | -34.1943 | -9.675
5.5427 | -34.1946 | -9.6854
5.5488 | -34.1949 | -9.6957
5.5549 | -34.1951 | -9.706
5.561 | -34.1954 | -9.7165
5.5671 | -34.1958 | -9.7266
5.5732 | -34.196 | -9.7371
5.5793 | -34.1964 | -9.7476
5.5854 | -34.1967 | -9.7578
5.5915 | -34.197 | -9.7684
5.5976 | -34.1972 | -9.7787
5.6037 | -34.1975 | -9.7891
5.6098 | -34.1978 | -9.7993
5.6159 | -34.1981 | -9.8094
5.622 | -34.1984 | -9.8198
5.6281 | -34.1987 | -9.8301
5.6342 | -34.199 | -9.8407
5.6403 | -34.1993 | -9.851
5.6464 | -34.1996 | -9.8614
5.6525 | -34.1998 | -9.8718
5.6586 | -34.2001 | -9.8818
5.6648 | -34.2005 | -9.8921
5.6709 | -34.2008 | -9.9026
5.677 | -34.2011 | -9.9129
5.6831 | -34.2014 | -9.9233
5.6892 | -34.2016 | -9.9337
5.6953 | -34.2019 | -9.944
5.7014 | -34.2022 | -9.9543
5.7075 | -34.2025 | -9.9646
5.7136 | -34.2028 | -9.9748
5.7197 | -34.203 | -9.9851
5.7258 | -34.2033 | -9.9952
5.7319 | -34.2036 | -10.0055
5.738 | -34.2039 | -10.0157
5.7441 | -34.2042 | -10.026
5.7502 | -34.2045 | -10.036
5.7563 | -34.2048 | -10.0463
5.7624 | -34.2052 | -10.0567
5.7685 | -34.2054 | -10.067
5.7746 | -34.2057 | -10.0776
5.7807 | -34.206 | -10.0879
5.7868 | -34.2062 | -10.0984
5.7929 | -34.2065 | -10.1086
5.799 | -34.2068 | -10.1186
5.8052 | -34.2071 | -10.1288
5.8113 | -34.2074 | -10.1389
5.8174 | -34.2076 | -10.1491
5.8235 | -34.2079 | -10.1593
5.8296 | -34.2082 | -10.1697
5.8357 | -34.2085 | -10.1799
5.8418 | -34.2088 | -10.1903
5.8479 | -34.209 | -10.2006
5.854 | -34.2093 | -10.211
5.8601 | -34.2096 | -10.2212
5.8662 | -34.2099 | -10.2316
5.8723 | -34.2101 | -10.2417
5.8784 | -34.2104 | -10.2518
5.8845 | -34.2107 | -10.262
5.8906 | -34.211 | -10.2723
5.8967 | -34.2113 | -10.2824
5.9028 | -34.2116 | -10.2924
5.9089 | -34.2119 | -10.3026
5.915 | -34.2122 | -10.3126
5.9211 | -34.2125 | -10.3228
5.9272 | -34.2128 | -10.3327
5.9333 | -34.2131 | -10.3428
5.9394 | -34.2134 | -10.353
5.9455 | -34.2137 | -10.363
5.9517 | -34.214 | -10.3731
5.9578 | -34.2143 | -10.3834
5.9639 | -34.2146 | -10.3933
5.97 | -34.215 | -10.4036
5.9761 | -34.2153 | -10.4137
5.9822 | -34.2155 | -10.424
5.9883 | -34.2158 | -10.4342
5.9944 | -34.2161 | -10.4445
6.0005 | -34.2165 | -10.4549
6.0066 | -34.2167 | -10.4649
6.0127 | -34.2171 | -10.4748
6.0188 | -34.2174 | -10.4852
6.0249 | -34.2176 | -10.4955
6.031 | -34.2179 | -10.5055
6.0371 | -34.2182 | -10.5157
6.0432 | -34.2185 | -10.5258
6.0493 | -34.2188 | -10.5358
6.0554 | -34.2191 | -10.5458
6.0615 | -34.2194 | -10.556
6.0676 | -34.2198 | -10.5664
6.0737 | -34.2201 | -10.5767
6.0798 | -34.2203 | -10.5868
6.0859 | -34.2207 | -10.5969
6.0921 | -34.221 | -10.6069
6.0982 | -34.2213 | -10.617
6.1043 | -34.2216 | -10.6271
6.1104 | -34.2219 | -10.6374
6.1165 | -34.2222 | -10.6472
6.1226 | -34.2225 | -10.6572
6.1287 | -34.2228 | -10.6673
6.1348 | -34.2231 | -10.6775
6.1409 | -34.2234 | -10.6875
6.147 | -34.2237 | -10.6976
6.1531 | -34.224 | -10.7077
6.1592 | -34.2243 | -10.7178
6.1653 | -34.2246 | -10.7278
6.1714 | -34.2249 | -10.738
6.1775 | -34.2252 | -10.7482
6.1836 | -34.2255 | -10.7583
6.1897 | -34.2258 | -10.7685
6.1958 | -34.2261 | -10.7788
6.2019 | -34.2265 | -10.789
6.208 | -34.2268 | -10.7991
6.2141 | -34.2271 | -10.8092
6.2202 | -34.2274 | -10.8194
6.2263 | -34.2277 | -10.8295
6.2325 | -34.228 | -10.8394
6.2386 | -34.2283 | -10.8492
6.2447 | -34.2287 | -10.8596
6.2508 | -34.229 | -10.8697
6.2569 | -34.2293 | -10.8798
6.263 | -34.2296 | -10.8898
6.2691 | -34.2299 | -10.8999
6.2752 | -34.2302 | -10.91
6.2813 | -34.2305 | -10.9201
6.2874 | -34.2308 | -10.9301
6.2935 | -34.2311 | -10.9402
6.2996 | -34.2313 | -10.9502
6.3057 | -34.2317 | -10.9604
6.3118 | -34.2319 | -10.9706
6.3179 | -34.2322 | -10.9807
6.324 | -34.2325 | -10.9908
6.3301 | -34.2328 | -11.0009
6.3362 | -34.2332 | -11.0109
6.3423 | -34.2335 | -11.0212
6.3484 | -34.2338 | -11.0313
6.3545 | -34.2341 | -11.0412
6.3606 | -34.2344 | -11.051
6.3667 | -34.2346 | -11.061
6.3728 | -34.235 | -11.0711
6.379 | -34.2353 | -11.0812
6.3851 | -34.2356 | -11.0911
6.3912 | -34.2359 | -11.1011
6.3973 | -34.2362 | -11.1112
6.4034 | -34.2364 | -11.1212
6.4095 | -34.2367 | -11.1312
6.4156 | -34.237 | -11.1413
6.4217 | -34.2373 | -11.1513
6.4278 | -34.2376 | -11.1614
6.4339 | -34.238 | -11.1714
6.44 | -34.2383 | -11.1814
6.4461 | -34.2386 | -11.1916
6.4522 | -34.2389 | -11.2015
6.4583 | -34.2391 | -11.2117
6.4644 | -34.2394 | -11.2218
6.4705 | -34.2397 | -11.2316
6.4766 | -34.2401 | -11.2417
6.4827 | -34.2404 | -11.2518
6.4888 | -34.2406 | -11.2617
6.4949 | -34.2409 | -11.2715
6.501 | -34.2412 | -11.2816
6.5071 | -34.2416 | -11.2915
6.5132 | -34.2419 | -11.3016
6.5194 | -34.2421 | -11.3117
6.5255 | -34.2425 | -11.3215
6.5316 | -34.2427 | -11.3317
6.5377 | -34.243 | -11.3417
6.5438 | -34.2433 | -11.3516
6.5499 | -34.2436 | -11.3616
6.556 | -34.2439 | -11.3715
6.5621 | -34.2442 | -11.3818
6.5682 | -34.2445 | -11.3919
6.5743 | -34.2449 | -11.4019
6.5804 | -34.2452 | -11.412
6.5865 | -34.2455 | -11.4221
6.5926 | -34.2458 | -11.432
6.5987 | -34.2461 | -11.4418
6.6048 | -34.2464 | -11.4519
6.6109 | -34.2467 | -11.462
6.617 | -34.247 | -11.4722
6.6231 | -34.2473 | -11.4822
6.6292 | -34.2476 | -11.4922
6.6353 | -34.2479 | -11.5025
6.6414 | -34.2482 | -11.5124
6.6475 | -34.2485 | -11.5225
6.6536 | -34.2488 | -11.5325
6.6597 | -34.2491 | -11.5423
6.6659 | -34.2494 | -11.5523
6.672 | -34.2497 | -11.5622
6.6781 | -34.25 | -11.5723
6.6842 | -34.2503 | -11.5822
6.6903 | -34.2507 | -11.5921
6.6964 | -34.251 | -11.6021
6.7025 | -34.2513 | -11.6119
6.7086 | -34.2516 | -11.6219
6.7147 | -34.2519 | -11.632
6.7208 | -34.2522 | -11.6421
6.7269 | -34.2525 | -11.652
6.733 | -34.2527 | -11.6618
6.7391 | -34.2531 | -11.6719
6.7452 | -34.2534 | -11.6816
6.7513 | -34.2537 | -11.6914
6.7574 | -34.254 | -11.7015
6.7635 | -34.2543 | -11.7113
6.7696 | -34.2546 | -11.7212
6.7757 | -34.255 | -11.7312
6.7818 | -34.2553 | -11.7412
6.7879 | -34.2556 | -11.751
6.794 | -34.2559 | -11.7608
6.8001 | -34.2562 | -11.7708
6.8063 | -34.2565 | -11.7808
6.8124 | -34.2568 | -11.7909
6.8185 | -34.2571 | -11.801
6.8246 | -34.2574 | -11.8109
6.8307 | -34.2577 | -11.8208
6.8368 | -34.258 | -11.831
6.8429 | -34.2583 | -11.841
6.849 | -34.2586 | -11.851
6.8551 | -34.2589 | -11.8608
6.8612 | -34.2592 | -11.8708
6.8673 | -34.2595 | -11.8807
6.8734 | -34.2598 | -11.8909
6.8795 | -34.2602 | -11.9008
6.8856 | -34.2605 | -11.9107
6.8917 | -34.2608 | -11.9207
6.8978 | -34.2611 | -11.9304
6.9039 | -34.2614 | -11.9401
6.91 | -34.2617 | -11.9501
6.9161 | -34.262 | -11.9599
6.9222 | -34.2623 | -11.9698
6.9283 | -34.2627 | -11.9795
6.9344 | -34.263 | -11.9894
6.9405 | -34.2633 | -11.9992
6.9466 | -34.2637 | -12.0094
6.9528 | -34.264 | -12.0195
6.9589 | -34.2643 | -12.0295
6.965 | -34.2646 | -12.0396
6.9711 | -34.265 | -12.0496
6.9772 | -34.2653 | -12.0595
6.9833 | -34.2656 | -12.0694
6.9894 | -34.266 | -12.0795
6.9955 | -34.2663 | -12.0893
7.0016 | -34.2666 | -12.0992
7.0077 | -34.2669 | -12.1091
7.0138 | -34.2673 | -12.1189
7.0199 | -34.2675 | -12.1288
7.026 | -34.2679 | -12.1386
7.0321 | -34.2682 | -12.1484
7.0382 | -34.2685 | -12.158
7.0443 | -34.2689 | -12.168
7.0504 | -34.2692 | -12.178
7.0565 | -34.2695 | -12.1879
7.0626 | -34.2698 | -12.1976
7.0687 | -34.2701 | -12.2071
7.0748 | -34.2705 | -12.2167
7.0809 | -34.2708 | -12.2266
7.087 | -34.2711 | -12.2366
7.0932 | -34.2714 | -12.2465
7.0993 | -34.2717 | -12.2564
7.1054 | -34.2721 | -12.2661
7.1115 | -34.2724 | -12.2758
7.1176 | -34.2727 | -12.2855
7.1237 | -34.273 | -12.2954
7.1298 | -34.2734 | -12.3051
7.1359 | -34.2737 | -12.315
7.142 | -34.274 | -12.3246
7.1481 | -34.2743 | -12.3343
7.1542 | -34.2746 | -12.3442
7.1603 | -34.2749 | -12.354
7.1664 | -34.2753 | -12.3638
7.1725 | -34.2756 | -12.3736
7.1786 | -34.276 | -12.3834
7.1847 | -34.2763 | -12.3933
7.1908 | -34.2766 | -12.4033
7.1969 | -34.2769 | -12.4133
7.203 | -34.2772 | -12.4229
7.2091 | -34.2776 | -12.4326
7.2152 | -34.2779 | -12.4424
7.2213 | -34.2782 | -12.452
7.2274 | -34.2786 | -12.4618
7.2335 | -34.2789 | -12.4715
7.2397 | -34.2792 | -12.4813
7.2458 | -34.2796 | -12.4908
7.2519 | -34.2799 | -12.5006
7.258 | -34.2802 | -12.5106
7.2641 | -34.2806 | -12.5204
7.2702 | -34.2809 | -12.5303
7.2763 | -34.2812 | -12.5402
7.2824 | -34.2816 | -12.55
7.2885 | -34.2819 | -12.5598
7.2946 | -34.2822 | -12.5696
7.3007 | -34.2825 | -12.5795
7.3068 | -34.2828 | -12.5894
7.3129 | -34.2832 | -12.5993
7.319 | -34.2835 | -12.6091
7.3251 | -34.2839 | -12.6187
7.3312 | -34.2842 | -12.6284
7.3373 | -34.2845 | -12.638
7.3434 | -34.2849 | -12.6477
7.3495 | -34.2852 | -12.6573
7.3556 | -34.2856 | -12.6672
7.3617 | -34.2859 | -12.6771
7.3678 | -34.2862 | -12.6868
7.3739 | -34.2865 | -12.6967
7.3801 | -34.2869 | -12.7067
7.3862 | -34.2872 | -12.7164
7.3923 | -34.2876 | -12.7264
7.3984 | -34.2879 | -12.7359
7.4045 | -34.2882 | -12.7456
7.4106 | -34.2886 | -12.7554
7.4167 | -34.2889 | -12.7652
7.4228 | -34.2893 | -12.7751
7.4289 | -34.2896 | -12.7848
7.435 | -34.29 | -12.7943
7.4411 | -34.2903 | -12.8041
7.4472 | -34.2906 | -12.8139
7.4533 | -34.291 | -12.8234
7.4594 | -34.2913 | -12.8334
7.4655 | -34.2917 | -12.843
7.4716 | -34.292 | -12.8529
7.4777 | -34.2923 | -12.8627
7.4838 | -34.2926 | -12.8722
7.4899 | -34.293 | -12.8818
7.496 | -34.2933 | -12.8914
7.5021 | -34.2936 | -12.9011
7.5082 | -34.294 | -12.9108
7.5143 | -34.2943 | -12.9206
7.5204 | -34.2946 | -12.9305
7.5266 | -34.295 | -12.94
7.5327 | -34.2953 | -12.9497
7.5388 | -34.2956 | -12.9596
7.5449 | -34.296 | -12.9696
7.551 | -34.2963 | -12.9792
7.5571 | -34.2967 | -12.989
7.5632 | -34.297 | -12.9988
7.5693 | -34.2974 | -13.0084
7.5754 | -34.2977 | -13.0179
7.5815 | -34.2981 | -13.0277
7.5876 | -34.2984 | -13.0376
7.5937 | -34.2988 | -13.0475
7.5998 | -34.2991 | -13.0571
7.6059 | -34.2995 | -13.0671
7.612 | -34.2998 | -13.0768
7.6181 | -34.3002 | -13.0864
7.6242 | -34.3005 | -13.0964
7.6303 | -34.3009 | -13.1063
7.6364 | -34.3012 | -13.1162
7.6425 | -34.3015 | -13.126
7.6486 | -34.3018 | -13.1357
7.6547 | -34.3022 | -13.1453
7.6608 | -34.3025 | -13.155
7.667 | -34.3028 | -13.1648
7.6731 | -34.3032 | -13.1746
7.6792 | -34.3035 | -13.1846
7.6853 | -34.3039 | -13.1945
7.6914 | -34.3042 | -13.2043
7.6975 | -34.3046 | -13.2141
7.7036 | -34.3049 | -13.2239
7.7097 | -34.3053 | -13.2338
7.7158 | -34.3056 | -13.2436
7.7219 | -34.306 | -13.2531
7.728 | -34.3063 | -13.263
7.7341 | -34.3067 | -13.2728
7.7402 | -34.307 | -13.2826
7.7463 | -34.3074 | -13.2924
7.7524 | -34.3077 | -13.3024
7.7585 | -34.308 | -13.3122
7.7646 | -34.3084 | -13.3219
7.7707 | -34.3087 | -13.3318
7.7768 | -34.309 | -13.3417
7.7829 | -34.3094 | -13.3517
7.789 | -34.3097 | -13.3617
7.7951 | -34.3101 | -13.3715
7.8012 | -34.3104 | -13.3813
7.8073 | -34.3108 | -13.3912
7.8135 | -34.3111 | -13.401
7.8196 | -34.3115 | -13.4109
7.8257 | -34.3118 | -13.4209
7.8318 | -34.3121 | -13.4308
7.8379 | -34.3125 | -13.4404
7.844 | -34.3128 | -13.4502
7.8501 | -34.3132 | -13.4601
7.8562 | -34.3135 | -13.4702
7.8623 | -34.3139 | -13.4799
7.8684 | -34.3142 | -13.4897
7.8745 | -34.3145 | -13.4996
7.8806 | -34.3149 | -13.5095
7.8867 | -34.3152 | -13.5191
7.8928 | -34.3156 | -13.5287
7.8989 | -34.316 | -13.5386
7.905 | -34.3163 | -13.5485
7.9111 | -34.3166 | -13.5583
7.9172 | -34.317 | -13.5681
7.9233 | -34.3173 | -13.578
7.9294 | -34.3176 | -13.5879
7.9355 | -34.318 | -13.5975
7.9416 | -34.3183 | -13.6074
7.9477 | -34.3186 | -13.6171
7.9539 | -34.319 | -13.6268
7.96 | -34.3193 | -13.6366
7.9661 | -34.3197 | -13.6465
7.9722 | -34.32 | -13.6561
7.9783 | -34.3204 | -13.6661
7.9844 | -34.3207 | -13.6758
7.9905 | -34.3211 | -13.6858
7.9966 | -34.3214 | -13.6958
8.0027 | -34.3217 | -13.7055
8.0088 | -34.3221 | -13.7153
8.0149 | -34.3224 | -13.7252
8.021 | -34.3227 | -13.7351
8.0271 | -34.3231 | -13.7451
8.0332 | -34.3234 | -13.7548
8.0393 | -34.3238 | -13.7647
8.0454 | -34.3241 | -13.7748
8.0515 | -34.3245 | -13.7846
8.0576 | -34.3249 | -13.7946
8.0637 | -34.3252 | -13.8043
8.0698 | -34.3255 | -13.814
8.0759 | -34.3258 | -13.8238
8.082 | -34.3262 | -13.8335
8.0881 | -34.3265 | -13.8434
8.0942 | -34.3269 | -13.8533
8.1004 | -34.3273 | -13.8632
8.1065 | -34.3276 | -13.8729
8.1126 | -34.3279 | -13.8828
8.1187 | -34.3283 | -13.8927
8.1248 | -34.3286 | -13.9023
8.1309 | -34.329 | -13.912
8.137 | -34.3293 | -13.9216
8.1431 | -34.3297 | -13.9317
8.1492 | -34.33 | -13.9415
8.1553 | -34.3303 | -13.9513
8.1614 | -34.3307 | -13.9613
8.1675 | -34.331 | -13.9712
8.1736 | -34.3313 | -13.9811
8.1797 | -34.3317 | -13.991
8.1858 | -34.332 | -14.0009
8.1919 | -34.3324 | -14.0107
8.198 | -34.3327 | -14.0203
8.2041 | -34.3331 | -14.0302
8.2102 | -34.3334 | -14.0401
8.2163 | -34.3338 | -14.0497
8.2224 | -34.3341 | -14.0596
8.2285 | -34.3344 | -14.0691
8.2346 | -34.3348 | -14.0789
8.2408 | -34.3351 | -14.0886
8.2469 | -34.3355 | -14.0985
8.253 | -34.3358 | -14.1082
8.2591 | -34.3362 | -14.1181
8.2652 | -34.3366 | -14.1279
8.2713 | -34.3369 | -14.1378
8.2774 | -34.3373 | -14.1477
8.2835 | -34.3376 | -14.1577
8.2896 | -34.3379 | -14.1674
8.2957 | -34.3383 | -14.1773
8.3018 | -34.3386 | -14.1871
8.3079 | -34.339 | -14.1968
8.314 | -34.3394 | -14.2067
8.3201 | -34.3397 | -14.2164
8.3262 | -34.34 | -14.2263
8.3323 | -34.3404 | -14.236
8.3384 | -34.3408 | -14.2459
8.3445 | -34.3411 | -14.2556
8.3506 | -34.3414 | -14.2655
8.3567 | -34.3418 | -14.2753
8.3628 | -34.3421 | -14.2852
8.3689 | -34.3425 | -14.2949
8.375 | -34.3428 | -14.3047
8.3812 | -34.3431 | -14.3143
8.3873 | -34.3435 | -14.3242
8.3934 | -34.3439 | -14.334
8.3995 | -34.3442 | -14.3437
8.4056 | -34.3445 | -14.3535
8.4117 | -34.3449 | -14.3634
8.4178 | -34.3452 | -14.3732
8.4239 | -34.3456 | -14.3828
8.43 | -34.3459 | -14.3927
8.4361 | -34.3462 | -14.4025
8.4422 | -34.3466 | -14.4122
8.4483 | -34.3469 | -14.4219
8.4544 | -34.3473 | -14.4318
8.4605 | -34.3476 | -14.4414
8.4666 | -34.348 | -14.4513
8.4727 | -34.3483 | -14.4612
8.4788 | -34.3486 | -14.4709
8.4849 | -34.349 | -14.4807
8.491 | -34.3493 | -14.4905
8.4971 | -34.3497 | -14.5001
8.5032 | -34.35 | -14.5098
8.5093 | -34.3504 | -14.5194
8.5154 | -34.3508 | -14.5293
8.5215 | -34.3511 | -14.5389
8.5277 | -34.3514 | -14.5487
8.5338 | -34.3518 | -14.5584
8.5399 | -34.3521 | -14.5682
8.546 | -34.3525 | -14.5778
8.5521 | -34.3529 | -14.5875
8.5582 | -34.3532 | -14.5973
8.5643 | -34.3535 | -14.6068
8.5704 | -34.3539 | -14.6164
8.5765 | -34.3542 | -14.6263
8.5826 | -34.3546 | -14.6357
8.5887 | -34.355 | -14.6454
8.5948 | -34.3553 | -14.6551
8.6009 | -34.3557 | -14.6647
8.607 | -34.356 | -14.6745
8.6131 | -34.3564 | -14.6843
8.6192 | -34.3568 | -14.6939
8.6253 | -34.3571 | -14.7035
8.6314 | -34.3575 | -14.7132
8.6375 | -34.3578 | -14.723
8.6436 | -34.3582 | -14.7327
8.6497 | -34.3585 | -14.7423
8.6558 | -34.3588 | -14.752
8.6619 | -34.3592 | -14.7618
8.6681 | -34.3595 | -14.7716
8.6742 | -34.3599 | -14.7813
8.6803 | -34.3603 | -14.7909
8.6864 | -34.3606 | -14.8004
8.6925 | -34.361 | -14.8099
8.6986 | -34.3613 | -14.8197
8.7047 | -34.3617 | -14.8295
8.7108 | -34.362 | -14.8392
8.7169 | -34.3624 | -14.8489
8.723 | -34.3627 | -14.8585
8.7291 | -34.3631 | -14.868
8.7352 | -34.3634 | -14.8779
8.7413 | -34.3638 | -14.8876
8.7474 | -34.3641 | -14.8972
8.7535 | -34.3645 | -14.9069
8.7596 | -34.3649 | -14.9166
8.7657 | -34.3652 | -14.9264
8.7718 | -34.3656 | -14.9361
8.7779 | -34.3659 | -14.9459
8.784 | -34.3663 | -14.9556
8.7901 | -34.3667 | -14.9653
8.7962 | -34.367 | -14.9748
8.8023 | -34.3674 | -14.9844
8.8084 | -34.3677 | -14.994
8.8146 | -34.3681 | -15.0037
8.8207 | -34.3684 | -15.0133
8.8268 | -34.3688 | -15.0227
8.8329 | -34.3692 | -15.0324
8.839 | -34.3695 | -15.042
8.8451 | -34.3699 | -15.0516
8.8512 | -34.3702 | -15.0611
8.8573 | -34.3706 | -15.0708
8.8634 | -34.371 | -15.0803
8.8695 | -34.3713 | -15.0901
8.8756 | -34.3717 | -15.0998
8.8817 | -34.372 | -15.1093
8.8878 | -34.3724 | -15.1192
8.8939 | -34.3727 | -15.1288
8.9 | -34.3731 | -15.1384
8.9061 | -34.3735 | -15.148
8.9122 | -34.3738 | -15.1576
8.9183 | -34.3742 | -15.1673
8.9244 | -34.3746 | -15.1771
8.9305 | -34.375 | -15.1865
8.9366 | -34.3754 | -15.1961
8.9427 | -34.3757 | -15.2058
8.9488 | -34.3761 | -15.2154
8.955 | -34.3764 | -15.2249
8.9611 | -34.3768 | -15.2347
8.9672 | -34.3772 | -15.2443
8.9733 | -34.3776 | -15.2538
8.9794 | -34.378 | -15.2633
8.9855 | -34.3784 | -15.273
8.9916 | -34.3787 | -15.2828
8.9977 | -34.3791 | -15.2923
9.0038 | -34.3795 | -15.3018
9.0099 | -34.3799 | -15.3115
9.016 | -34.3802 | -15.3211
9.0221 | -34.3806 | -15.3305
9.0282 | -34.381 | -15.34
9.0343 | -34.3813 | -15.3495
9.0404 | -34.3817 | -15.3591
9.0465 | -34.3821 | -15.3689
9.0526 | -34.3825 | -15.3784
9.0587 | -34.3829 | -15.388
9.0648 | -34.3833 | -15.3975
9.0709 | -34.3836 | -15.407
9.077 | -34.384 | -15.4167
9.0831 | -34.3844 | -15.4264
9.0892 | -34.3848 | -15.436
9.0953 | -34.3852 | -15.4457
9.1015 | -34.3855 | -15.4553
9.1076 | -34.3859 | -15.4647
9.1137 | -34.3863 | -15.4743
9.1198 | -34.3867 | -15.4838
9.1259 | -34.3871 | -15.4934
9.132 | -34.3875 | -15.5031
9.1381 | -34.3879 | -15.5126
9.1442 | -34.3883 | -15.5219
9.1503 | -34.3886 | -15.5315
9.1564 | -34.389 | -15.5411
9.1625 | -34.3894 | -15.5507
9.1686 | -34.3898 | -15.5602
9.1747 | -34.3902 | -15.5698
9.1808 | -34.3906 | -15.5794
9.1869 | -34.391 | -15.5889
9.193 | -34.3914 | -15.5985
9.1991 | -34.3917 | -15.6081
9.2052 | -34.3921 | -15.6177
9.2113 | -34.3925 | -15.6273
9.2174 | -34.3929 | -15.6368
9.2235 | -34.3933 | -15.6464
9.2296 | -34.3937 | -15.6559
9.2357 | -34.3941 | -15.6656
9.2419 | -34.3945 | -15.6751
9.248 | -34.3949 | -15.6847
9.2541 | -34.3953 | -15.6943
9.2602 | -34.3957 | -15.704
9.2663 | -34.3961 | -15.7133
9.2724 | -34.3965 | -15.7228
9.2785 | -34.3969 | -15.7323
9.2846 | -34.3973 | -15.7418
9.2907 | -34.3977 | -15.7513
9.2968 | -34.3981 | -15.761
9.3029 | -34.3985 | -15.7706
9.309 | -34.3989 | -15.7802
9.3151 | -34.3992 | -15.7896
9.3212 | -34.3997 | -15.799
9.3273 | -34.4001 | -15.8085
9.3334 | -34.4005 | -15.8179
9.3395 | -34.4009 | -15.8273
9.3456 | -34.4012 | -15.8367
9.3517 | -34.4016 | -15.8463
9.3578 | -34.402 | -15.8558
9.3639 | -34.4024 | -15.8652
9.37 | -34.4028 | -15.8747
9.3761 | -34.4032 | -15.8844
9.3822 | -34.4037 | -15.894
9.3884 | -34.4041 | -15.9033
9.3945 | -34.4045 | -15.9128
9.4006 | -34.4049 | -15.9223
9.4067 | -34.4053 | -15.9319
9.4128 | -34.4057 | -15.9415
9.4189 | -34.4061 | -15.9511
9.425 | -34.4065 | -15.9606
9.4311 | -34.4069 | -15.97
9.4372 | -34.4073 | -15.9795
9.4433 | -34.4077 | -15.989
9.4494 | -34.4081 | -15.9984
9.4555 | -34.4085 | -16.008
9.4616 | -34.409 | -16.0175
9.4677 | -34.4094 | -16.027
9.4738 | -34.4098 | -16.0365
9.4799 | -34.4102 | -16.046
9.486 | -34.4106 | -16.0555
9.4921 | -34.411 | -16.065
9.4982 | -34.4114 | -16.0743
9.5043 | -34.4118 | -16.0838
9.5104 | -34.4122 | -16.0932
9.5165 | -34.4126 | -16.1027
9.5226 | -34.413 | -16.1123
9.5288 | -34.4134 | -16.1217
9.5349 | -34.4138 | -16.1312
9.541 | -34.4142 | -16.1405
9.5471 | -34.4146 | -16.15
9.5532 | -34.415 | -16.1595
9.5593 | -34.4154 | -16.1691
9.5654 | -34.4158 | -16.1786
9.5715 | -34.4162 | -16.188
9.5776 | -34.4166 | -16.1976
9.5837 | -34.417 | -16.207
9.5898 | -34.4174 | -16.2166
9.5959 | -34.4178 | -16.2261
9.602 | -34.4182 | -16.2354
9.6081 | -34.4186 | -16.2449
9.6142 | -34.4191 | -16.2543
9.6203 | -34.4194 | -16.2638
9.6264 | -34.4198 | -16.2731
9.6325 | -34.4202 | -16.2827
9.6386 | -34.4206 | -16.2921
9.6447 | -34.421 | -16.3016
9.6508 | -34.4214 | -16.311
9.6569 | -34.4218 | -16.3204
9.663 | -34.4223 | -16.33
9.6691 | -34.4227 | -16.3395
9.6753 | -34.4231 | -16.3488
9.6814 | -34.4235 | -16.3581
9.6875 | -34.4239 | -16.3674
9.6936 | -34.4243 | -16.3767
9.6997 | -34.4247 | -16.3863
9.7058 | -34.4252 | -16.3957
9.7119 | -34.4256 | -16.4052
9.718 | -34.426 | -16.4148
9.7241 | -34.4264 | -16.4241
9.7302 | -34.4268 | -16.4336
9.7363 | -34.4272 | -16.4429
9.7424 | -34.4277 | -16.4522
9.7485 | -34.428 | -16.4617
9.7546 | -34.4284 | -16.471
9.7607 | -34.4289 | -16.4806
9.7668 | -34.4293 | -16.4901
9.7729 | -34.4297 | -16.4995
9.779 | -34.4301 | -16.5087
9.7851 | -34.4305 | -16.5183
9.7912 | -34.431 | -16.5276
9.7973 | -34.4314 | -16.5371
9.8034 | -34.4318 | -16.5464
9.8095 | -34.4322 | -16.556
9.8157 | -34.4327 | -16.5654
9.8218 | -34.4331 | -16.575
9.8279 | -34.4335 | -16.5843
9.834 | -34.4339 | -16.5938
9.8401 | -34.4343 | -16.6032
9.8462 | -34.4347 | -16.6124
9.8523 | -34.4351 | -16.6218
9.8584 | -34.4355 | -16.6312
9.8645 | -34.4359 | -16.6408
9.8706 | -34.4363 | -16.6502
9.8767 | -34.4367 | -16.6596
9.8828 | -34.4371 | -16.6689
9.8889 | -34.4375 | -16.6781
9.895 | -34.4379 | -16.6876
9.9011 | -34.4384 | -16.697
9.9072 | -34.4388 | -16.7064
9.9133 | -34.4392 | -16.7157
9.9194 | -34.4396 | -16.725
9.9255 | -34.44 | -16.7344
9.9316 | -34.4404 | -16.7437
9.9377 | -34.4408 | -16.7531
9.9438 | -34.4412 | -16.7625
9.9499 | -34.4416 | -16.7721
9.956 | -34.442 | -16.7815
9.9622 | -34.4424 | -16.791
9.9683 | -34.4428 | -16.8004
9.9744 | -34.4432 | -16.8099
9.9805 | -34.4437 | -16.8193
9.9866 | -34.4441 | -16.8286
9.9927 | -34.4445 | -16.8379
9.9988 | -34.4449 | -16.8474
10.0049 | -34.4453 | -16.8568
10.011 | -34.4457 | -16.8662
10.0171 | -34.4461 | -16.8755
10.0232 | -34.4465 | -16.8847
10.0293 | -34.447 | -16.8942
10.0354 | -34.4474 | -16.9036
10.0415 | -34.4478 | -16.913
10.0476 | -34.4482 | -16.9225
10.0537 | -34.4486 | -16.932
10.0598 | -34.449 | -16.9413
10.0659 | -34.4494 | -16.9507
10.072 | -34.4498 | -16.9599
10.0781 | -34.4503 | -16.9692
10.0842 | -34.4507 | -16.9786
10.0903 | -34.4511 | -16.988
10.0964 | -34.4515 | -16.9971
10.1026 | -34.4519 | -17.0065
10.1087 | -34.4523 | -17.0158
10.1148 | -34.4527 | -17.0252
10.1209 | -34.4531 | -17.0345
10.127 | -34.4535 | -17.0437
10.1331 | -34.4539 | -17.0529
10.1392 | -34.4543 | -17.0621
10.1453 | -34.4547 | -17.0714
10.1514 | -34.4551 | -17.0808
10.1575 | -34.4555 | -17.0902
10.1636 | -34.456 | -17.0993
10.1697 | -34.4564 | -17.1088
10.1758 | -34.4568 | -17.118
10.1819 | -34.4572 | -17.1272
10.188 | -34.4576 | -17.1366
10.1941 | -34.458 | -17.1459
10.2002 | -34.4584 | -17.1553
10.2063 | -34.4588 | -17.1646
10.2124 | -34.4592 | -17.174
10.2185 | -34.4596 | -17.1833
10.2246 | -34.46 | -17.1925
10.2307 | -34.4604 | -17.2019
10.2368 | -34.4608 | -17.2109
10.2429 | -34.4612 | -17.2203
10.2491 | -34.4616 | -17.2297
10.2552 | -34.462 | -17.239
10.2613 | -34.4625 | -17.2484
10.2674 | -34.4629 | -17.2577
10.2735 | -34.4633 | -17.2672
10.2796 | -34.4637 | -17.2765
10.2857 | -34.4641 | -17.2859
10.2918 | -34.4645 | -17.2954
10.2979 | -34.4649 | -17.3047
10.304 | -34.4654 | -17.3138
10.3101 | -34.4657 | -17.3232
10.3162 | -34.4661 | -17.3326
10.3223 | -34.4666 | -17.3418
10.3284 | -34.467 | -17.3511
10.3345 | -34.4674 | -17.3603
10.3406 | -34.4678 | -17.3695
10.3467 | -34.4682 | -17.3788
10.3528 | -34.4686 | -17.388
10.3589 | -34.469 | -17.3973
10.365 | -34.4694 | -17.4067
10.3711 | -34.4698 | -17.4161
10.3772 | -34.4702 | -17.4254
10.3833 | -34.4706 | -17.4347
10.3895 | -34.4711 | -17.444
10.3956 | -34.4715 | -17.4533
10.4017 | -34.4719 | -17.4624
10.4078 | -34.4723 | -17.4717
10.4139 | -34.4727 | -17.481
10.42 | -34.4731 | -17.4904
10.4261 | -34.4735 | -17.4997
10.4322 | -34.4739 | -17.5089
10.4383 | -34.4743 | -17.5182
10.4444 | -34.4747 | -17.5273
10.4505 | -34.4751 | -17.5367
10.4566 | -34.4755 | -17.5459
10.4627 | -34.4759 | -17.555
10.4688 | -34.4763 | -17.5642
10.4749 | -34.4767 | -17.5735
10.481 | -34.4771 | -17.5828
10.4871 | -34.4775 | -17.592
10.4932 | -34.4779 | -17.6011
10.4993 | -34.4783 | -17.6105
10.5054 | -34.4788 | -17.6197
10.5115 | -34.4792 | -17.629
10.5176 | -34.4796 | -17.6381
10.5237 | -34.48 | -17.6475
10.5298 | -34.4804 | -17.6567
10.536 | -34.4808 | -17.666
10.5421 | -34.4812 | -17.6751
10.5482 | -34.4816 | -17.6844
10.5543 | -34.4821 | -17.6935
10.5604 | -34.4825 | -17.7026
10.5665 | -34.4829 | -17.7117
10.5726 | -34.4833 | -17.721
10.5787 | -34.4837 | -17.7302
10.5848 | -34.4841 | -17.7391
10.5909 | -34.4845 | -17.7484
10.597 | -34.4849 | -17.7577
10.6031 | -34.4853 | -17.7671
10.6092 | -34.4857 | -17.7763
10.6153 | -34.4861 | -17.7856
10.6214 | -34.4865 | -17.7946
10.6275 | -34.4869 | -17.8038
10.6336 | -34.4873 | -17.8131
10.6397 | -34.4877 | -17.8223
10.6458 | -34.4881 | -17.8315
10.6519 | -34.4885 | -17.8408
10.658 | -34.4889 | -17.8499
10.6641 | -34.4894 | -17.8593
10.6702 | -34.4898 | -17.8685
10.6764 | -34.4902 | -17.8778
10.6825 | -34.4906 | -17.8869
10.6886 | -34.491 | -17.8959
10.6947 | -34.4914 | -17.9051
10.7008 | -34.4918 | -17.9143
10.7069 | -34.4922 | -17.9236
10.713 | -34.4926 | -17.9328
10.7191 | -34.4931 | -17.942
10.7252 | -34.4935 | -17.9509
10.7313 | -34.4939 | -17.96
10.7374 | -34.4943 | -17.9692
10.7435 | -34.4947 | -17.9784
10.7496 | -34.4951 | -17.9875
10.7557 | -34.4955 | -17.9968
10.7618 | -34.4959 | -18.0057
10.7679 | -34.4963 | -18.0149
10.774 | -34.4967 | -18.0241
10.7801 | -34.4971 | -18.0333
10.7862 | -34.4975 | -18.0423
10.7923 | -34.4979 | -18.0516
10.7984 | -34.4983 | -18.0606
10.8045 | -34.4987 | -18.0697
10.8106 | -34.4991 | -18.0789
10.8168 | -34.4995 | -18.088
10.8229 | -34.5 | -18.0971
10.829 | -34.5003 | -18.1062
10.8351 | -34.5008 | -18.1151
10.8412 | -34.5012 | -18.1242
10.8473 | -34.5016 | -18.1332
10.8534 | -34.502 | -18.1423
10.8595 | -34.5024 | -18.1514
10.8656 | -34.5029 | -18.1603
10.8717 | -34.5033 | -18.1695
10.8778 | -34.5037 | -18.1786
10.8839 | -34.5041 | -18.1878
10.89 | -34.5045 | -18.1969
10.8961 | -34.5049 | -18.206
10.9022 | -34.5053 | -18.215
10.9083 | -34.5057 | -18.224
10.9144 | -34.5061 | -18.2331
10.9205 | -34.5065 | -18.2422
10.9266 | -34.507 | -18.2513
10.9327 | -34.5074 | -18.2603
10.9388 | -34.5078 | -18.2693
10.9449 | -34.5082 | -18.2784
10.951 | -34.5086 | -18.2875
10.9571 | -34.509 | -18.2966
10.9633 | -34.5094 | -18.3058
10.9694 | -34.5098 | -18.3151
10.9755 | -34.5103 | -18.3241
10.9816 | -34.5107 | -18.3333
10.9877 | -34.5111 | -18.3423
10.9938 | -34.5115 | -18.3515
10.9999 | -34.5119 | -18.3605
11.006 | -34.5123 | -18.3693
11.0121 | -34.5127 | -18.3784
11.0182 | -34.5132 | -18.3875
11.0243 | -34.5136 | -18.3965
11.0304 | -34.514 | -18.4056
11.0365 | -34.5144 | -18.4146
11.0426 | -34.5148 | -18.4235
11.0487 | -34.5153 | -18.4327
11.0548 | -34.5157 | -18.4416
11.0609 | -34.5161 | -18.4507
11.067 | -34.5165 | -18.4598
11.0731 | -34.5169 | -18.4688
11.0792 | -34.5174 | -18.4777
11.0853 | -34.5178 | -18.4868
11.0914 | -34.5182 | -18.4959
11.0975 | -34.5186 | -18.5049
11.1037 | -34.519 | -18.5139
11.1098 | -34.5194 | -18.5231
11.1159 | -34.5199 | -18.5319
11.122 | -34.5203 | -18.5409
11.1281 | -34.5207 | -18.55
11.1342 | -34.5211 | -18.559
11.1403 | -34.5215 | -18.5682
11.1464 | -34.5219 | -18.5772
11.1525 | -34.5224 | -18.586
11.1586 | -34.5228 | -18.5952
11.1647 | -34.5232 | -18.6041
11.1708 | -34.5236 | -18.6131
11.1769 | -34.524 | -18.6222
11.183 | -34.5244 | -18.631
11.1891 | -34.5249 | -18.64
11.1952 | -34.5253 | -18.649
11.2013 | -34.5257 | -18.6581
11.2074 | -34.5261 | -18.6671
11.2135 | -34.5265 | -18.6762
11.2196 | -34.527 | -18.6851
11.2257 | -34.5274 | -18.6942
11.2318 | -34.5278 | -18.7031
11.2379 | -34.5283 | -18.7122
11.244 | -34.5287 | -18.7213
11.2502 | -34.5291 | -18.7302
11.2563 | -34.5296 | -18.7392
11.2624 | -34.53 | -18.7481
11.2685 | -34.5304 | -18.757
11.2746 | -34.5308 | -18.7659
11.2807 | -34.5313 | -18.7751
11.2868 | -34.5317 | -18.7841
11.2929 | -34.5321 | -18.7928
11.299 | -34.5326 | -18.8019
11.3051 | -34.533 | -18.8111
11.3112 | -34.5334 | -18.8201
11.3173 | -34.5338 | -18.829
11.3234 | -34.5343 | -18.8377
11.3295 | -34.5347 | -18.8468
11.3356 | -34.5351 | -18.856
11.3417 | -34.5355 | -18.8648
11.3478 | -34.5359 | -18.8739
11.3539 | -34.5363 | -18.883
11.36 | -34.5368 | -18.8918
11.3661 | -34.5372 | -18.9007
11.3722 | -34.5376 | -18.9095
11.3783 | -34.538 | -18.9186
11.3844 | -34.5385 | -18.9275
11.3906 | -34.5389 | -18.9365
11.3967 | -34.5393 | -18.9453
11.4028 | -34.5398 | -18.9544
11.4089 | -34.5402 | -18.9632
11.415 | -34.5406 | -18.9721
11.4211 | -34.541 | -18.9812
11.4272 | -34.5414 | -18.9902
11.4333 | -34.5418 | -18.999
11.4394 | -34.5422 | -19.0079
11.4455 | -34.5426 | -19.0168
11.4516 | -34.5431 | -19.0257
11.4577 | -34.5435 | -19.0347
11.4638 | -34.5439 | -19.0434
11.4699 | -34.5444 | -19.0524
11.476 | -34.5448 | -19.0615
11.4821 | -34.5452 | -19.0704
11.4882 | -34.5456 | -19.0794
11.4943 | -34.5461 | -19.0884
11.5004 | -34.5465 | -19.0971
11.5065 | -34.5469 | -19.106
11.5126 | -34.5473 | -19.1148
11.5187 | -34.5478 | -19.1238
11.5248 | -34.5482 | -19.1328
11.5309 | -34.5486 | -19.1417
11.5371 | -34.5491 | -19.1505
11.5432 | -34.5495 | -19.1595
11.5493 | -34.5499 | -19.1685
11.5554 | -34.5503 | -19.1775
11.5615 | -34.5508 | -19.1865
11.5676 | -34.5512 | -19.1954
11.5737 | -34.5516 | -19.2041
11.5798 | -34.5521 | -19.2129
11.5859 | -34.5525 | -19.2218
11.592 | -34.5529 | -19.2306
11.5981 | -34.5533 | -19.2394
11.6042 | -34.5538 | -19.2483
11.6103 | -34.5542 | -19.2571
11.6164 | -34.5546 | -19.2659
11.6225 | -34.5551 | -19.2747
11.6286 | -34.5555 | -19.2836
11.6347 | -34.5559 | -19.2926
11.6408 | -34.5563 | -19.3013
11.6469 | -34.5568 | -19.3102
11.653 | -34.5572 | -19.319
11.6591 | -34.5576 | -19.3279
11.6652 | -34.558 | -19.3366
11.6713 | -34.5585 | -19.3454
11.6775 | -34.5589 | -19.3542
11.6836 | -34.5593 | -19.363
11.6897 | -34.5598 | -19.3719
11.6958 | -34.5602 | -19.3808
11.7019 | -34.5606 | -19.3898
11.708 | -34.5611 | -19.3987
11.7141 | -34.5615 | -19.4074
11.7202 | -34.562 | -19.4163
11.7263 | -34.5624 | -19.4252
11.7324 | -34.5628 | -19.4342
11.7385 | -34.5632 | -19.4431
11.7446 | -34.5636 | -19.4518
11.7507 | -34.5641 | -19.4604
11.7568 | -34.5645 | -19.4693
11.7629 | -34.5649 | -19.4782
11.769 | -34.5653 | -19.4869
11.7751 | -34.5658 | -19.4958
11.7812 | -34.5662 | -19.5044
11.7873 | -34.5666 | -19.5133
11.7934 | -34.5671 | -19.5221
11.7995 | -34.5675 | -19.5309
11.8056 | -34.568 | -19.5397
11.8117 | -34.5684 | -19.5485
11.8178 | -34.5688 | -19.5572
11.824 | -34.5693 | -19.5659
11.8301 | -34.5697 | -19.5747
11.8362 | -34.5701 | -19.5834
11.8423 | -34.5705 | -19.5922
11.8484 | -34.5709 | -19.6011
11.8545 | -34.5714 | -19.6097
11.8606 | -34.5718 | -19.6184
11.8667 | -34.5723 | -19.6273
11.8728 | -34.5727 | -19.636
11.8789 | -34.5731 | -19.6448
11.885 | -34.5736 | -19.6536
11.8911 | -34.574 | -19.6623
11.8972 | -34.5745 | -19.6711
11.9033 | -34.5749 | -19.6798
11.9094 | -34.5753 | -19.6886
11.9155 | -34.5757 | -19.6974
11.9216 | -34.5762 | -19.706
11.9277 | -34.5766 | -19.7147
11.9338 | -34.577 | -19.7235
11.9399 | -34.5775 | -19.7324
11.946 | -34.5779 | -19.7411
11.9521 | -34.5783 | -19.7498
11.9582 | -34.5788 | -19.7585
11.9644 | -34.5792 | -19.7672
11.9705 | -34.5797 | -19.7759
11.9766 | -34.5801 | -19.7846
11.9827 | -34.5806 | -19.7934
11.9888 | -34.581 | -19.8021
11.9949 | -34.5815 | -19.8105
12.001 | -34.5819 | -19.8193
12.0071 | -34.5823 | -19.8279
12.0132 | -34.5828 | -19.8366
12.0193 | -34.5832 | -19.8455
12.0254 | -34.5836 | -19.8542
12.0315 | -34.5841 | -19.8627
12.0376 | -34.5845 | -19.8713
12.0437 | -34.5849 | -19.8801
12.0498 | -34.5854 | -19.8889
12.0559 | -34.5858 | -19.8975
12.062 | -34.5863 | -19.9059
12.0681 | -34.5867 | -19.9147
12.0742 | -34.5872 | -19.9233
12.0803 | -34.5876 | -19.9321
12.0864 | -34.588 | -19.9409
12.0925 | -34.5885 | -19.9497
12.0986 | -34.5889 | -19.9583
12.1047 | -34.5894 | -19.967
12.1109 | -34.5898 | -19.9755
12.117 | -34.5902 | -19.9842
12.1231 | -34.5907 | -19.9928
12.1292 | -34.5911 | -20.0014
12.1353 | -34.5915 | -20.01
12.1414 | -34.592 | -20.0187
12.1475 | -34.5924 | -20.0275
12.1536 | -34.5929 | -20.0363
12.1597 | -34.5933 | -20.045
12.1658 | -34.5938 | -20.0536
12.1719 | -34.5942 | -20.062
12.178 | -34.5946 | -20.0707
12.1841 | -34.5951 | -20.0795
12.1902 | -34.5955 | -20.0882
12.1963 | -34.596 | -20.0969
12.2024 | -34.5964 | -20.1054
12.2085 | -34.5969 | -20.1141
12.2146 | -34.5973 | -20.1227
12.2207 | -34.5977 | -20.1314
12.2268 | -34.5982 | -20.1401
12.2329 | -34.5986 | -20.1487
12.239 | -34.5991 | -20.1571
12.2451 | -34.5995 | -20.1657
12.2513 | -34.6 | -20.1745
12.2574 | -34.6004 | -20.1831
12.2635 | -34.6009 | -20.1918
12.2696 | -34.6013 | -20.2004
12.2757 | -34.6018 | -20.2087
12.2818 | -34.6022 | -20.2174
12.2879 | -34.6027 | -20.2261
12.294 | -34.6031 | -20.2346
12.3001 | -34.6036 | -20.2432
12.3062 | -34.604 | -20.2518
12.3123 | -34.6045 | -20.2603
12.3184 | -34.6049 | -20.2687
12.3245 | -34.6054 | -20.2772
12.3306 | -34.6058 | -20.2858
12.3367 | -34.6062 | -20.2945
12.3428 | -34.6067 | -20.303
12.3489 | -34.6071 | -20.3114
12.355 | -34.6076 | -20.3201
12.3611 | -34.608 | -20.3288
12.3672 | -34.6085 | -20.3373
12.3733 | -34.6089 | -20.346
12.3794 | -34.6094 | -20.3543
12.3855 | -34.6098 | -20.3628
12.3916 | -34.6102 | -20.3714
12.3978 | -34.6107 | -20.38
12.4039 | -34.6112 | -20.3887
12.41 | -34.6116 | -20.3973
12.4161 | -34.6121 | -20.4056
12.4222 | -34.6125 | -20.4142
12.4283 | -34.6129 | -20.4227
12.4344 | -34.6134 | -20.4313
12.4405 | -34.6138 | -20.44
12.4466 | -34.6143 | -20.4485
12.4527 | -34.6147 | -20.4569
12.4588 | -34.6152 | -20.4655
12.4649 | -34.6156 | -20.4741
12.471 | -34.6161 | -20.4828
12.4771 | -34.6165 | -20.4913
12.4832 | -34.617 | -20.4998
12.4893 | -34.6174 | -20.5082
12.4954 | -34.6179 | -20.5166
12.5015 | -34.6183 | -20.5251
12.5076 | -34.6188 | -20.5337
12.5137 | -34.6192 | -20.5423
12.5198 | -34.6197 | -20.5507
12.5259 | -34.6201 | -20.5593
12.532 | -34.6205 | -20.5678
12.5382 | -34.621 | -20.5764
12.5443 | -34.6214 | -20.585
12.5504 | -34.6219 | -20.5936
12.5565 | -34.6224 | -20.6021
12.5626 | -34.6228 | -20.6107
12.5687 | -34.6233 | -20.6193
12.5748 | -34.6237 | -20.6279
12.5809 | -34.6241 | -20.6364
12.587 | -34.6246 | -20.6451
12.5931 | -34.6251 | -20.6534
12.5992 | -34.6255 | -20.662
12.6053 | -34.626 | -20.6705
12.6114 | -34.6264 | -20.6789
12.6175 | -34.6269 | -20.6875
12.6236 | -34.6273 | -20.696
12.6297 | -34.6278 | -20.7045
12.6358 | -34.6282 | -20.7129
12.6419 | -34.6287 | -20.7214
12.648 | -34.6291 | -20.73
12.6541 | -34.6296 | -20.7385
12.6602 | -34.63 | -20.7468
12.6663 | -34.6305 | -20.7554
12.6724 | -34.631 | -20.7638
12.6785 | -34.6314 | -20.7723
12.6847 | -34.6318 | -20.7806
12.6908 | -34.6323 | -20.7892
12.6969 | -34.6328 | -20.7976
12.703 | -34.6332 | -20.8062
12.7091 | -34.6337 | -20.8147
12.7152 | -34.6341 | -20.8231
12.7213 | -34.6346 | -20.8316
12.7274 | -34.635 | -20.8401
12.7335 | -34.6355 | -20.8484
12.7396 | -34.6359 | -20.8568
12.7457 | -34.6364 | -20.8653
12.7518 | -34.6368 | -20.8738
12.7579 | -34.6373 | -20.8824
12.764 | -34.6377 | -20.891
12.7701 | -34.6382 | -20.8995
12.7762 | -34.6386 | -20.9079
12.7823 | -34.6391 | -20.9164
12.7884 | -34.6395 | -20.925
12.7945 | -34.64 | -20.9336
12.8006 | -34.6405 | -20.942
12.8067 | -34.6409 | -20.9505
12.8128 | -34.6414 | -20.9588
12.8189 | -34.6418 | -20.9672
12.8251 | -34.6422 | -20.9757
12.8312 | -34.6427 | -20.9844
12.8373 | -34.6432 | -20.9928
12.8434 | -34.6436 | -21.0013
12.8495 | -34.6441 | -21.0097
12.8556 | -34.6445 | -21.0181
12.8617 | -34.645 | -21.0266
12.8678 | -34.6454 | -21.0352
12.8739 | -34.6459 | -21.0436
12.88 | -34.6463 | -21.0521
12.8861 | -34.6467 | -21.0605
12.8922 | -34.6472 | -21.069
12.8983 | -34.6477 | -21.0775
12.9044 | -34.6481 | -21.0858
12.9105 | -34.6486 | -21.0942
12.9166 | -34.649 | -21.1027
12.9227 | -34.6495 | -21.1112
12.9288 | -34.6499 | -21.1197
12.9349 | -34.6504 | -21.1282
12.941 | -34.6509 | -21.1365
12.9471 | -34.6513 | -21.1451
12.9532 | -34.6518 | -21.1535
12.9593 | -34.6522 | -21.1621
12.9654 | -34.6527 | -21.1707
12.9716 | -34.6532 | -21.1792
12.9777 | -34.6536 | -21.1877
12.9838 | -34.6541 | -21.1961
12.9899 | -34.6545 | -21.2045
12.996 | -34.6549 | -21.213
13.0021 | -34.6554 | -21.2215
13.0082 | -34.6559 | -21.2298
13.0143 | -34.6564 | -21.2381
13.0204 | -34.6568 | -21.2464
13.0265 | -34.6573 | -21.2548
13.0326 | -34.6577 | -21.2632
13.0387 | -34.6582 | -21.2718
13.0448 | -34.6586 | -21.2802
13.0509 | -34.6591 | -21.2887
13.057 | -34.6595 | -21.297
13.0631 | -34.66 | -21.3054
13.0692 | -34.6605 | -21.3137
13.0753 | -34.6609 | -21.3223
13.0814 | -34.6614 | -21.3309
13.0875 | -34.6619 | -21.3392
13.0936 | -34.6623 | -21.3477
13.0997 | -34.6628 | -21.356
13.1058 | -34.6632 | -21.3644
13.112 | -34.6637 | -21.3729
13.1181 | -34.6641 | -21.3812
13.1242 | -34.6646 | -21.3898
13.1303 | -34.6651 | -21.3982
13.1364 | -34.6655 | -21.4066
13.1425 | -34.6659 | -21.415
13.1486 | -34.6664 | -21.4234
13.1547 | -34.6668 | -21.4319
13.1608 | -34.6673 | -21.4402
13.1669 | -34.6677 | -21.4486
13.173 | -34.6682 | -21.4571
13.1791 | -34.6687 | -21.4655
13.1852 | -34.6691 | -21.4739
13.1913 | -34.6696 | -21.4823
13.1974 | -34.67 | -21.4908
13.2035 | -34.6705 | -21.4991
13.2096 | -34.671 | -21.5075
13.2157 | -34.6714 | -21.5157
13.2218 | -34.6719 | -21.5241
13.2279 | -34.6724 | -21.5324
13.234 | -34.6728 | -21.5407
13.2401 | -34.6733 | -21.5492
13.2462 | -34.6737 | -21.5576
13.2524 | -34.6742 | -21.5661
13.2585 | -34.6747 | -21.5743
13.2646 | -34.6751 | -21.5827
13.2707 | -34.6756 | -21.5912
13.2768 | -34.6761 | -21.5997
13.2829 | -34.6765 | -21.6083
13.289 | -34.677 | -21.6166
13.2951 | -34.6774 | -21.625
13.3012 | -34.6779 | -21.6334
13.3073 | -34.6783 | -21.642
13.3134 | -34.6788 | -21.6505
13.3195 | -34.6792 | -21.6588
13.3256 | -34.6796 | -21.667
13.3317 | -34.6801 | -21.6754
13.3378 | -34.6805 | -21.6838
13.3439 | -34.681 | -21.6923
13.35 | -34.6815 | -21.7007
13.3561 | -34.6819 | -21.7091
13.3622 | -34.6824 | -21.7174
13.3683 | -34.6829 | -21.7257
13.3744 | -34.6833 | -21.734
13.3805 | -34.6838 | -21.7425
13.3866 | -34.6842 | -21.7508
13.3927 | -34.6847 | -21.7592
13.3989 | -34.6851 | -21.7675
13.405 | -34.6856 | -21.7759
13.4111 | -34.686 | -21.7842
13.4172 | -34.6865 | -21.7926
13.4233 | -34.687 | -21.8011
13.4294 | -34.6874 | -21.8096
13.4355 | -34.6879 | -21.8177
13.4416 | -34.6883 | -21.826
13.4477 | -34.6888 | -21.8344
13.4538 | -34.6892 | -21.8428
13.4599 | -34.6897 | -21.8511
13.466 | -34.6902 | -21.8595
13.4721 | -34.6906 | -21.8678
13.4782 | -34.6911 | -21.8762
13.4843 | -34.6915 | -21.8846
13.4904 | -34.692 | -21.8928
13.4965 | -34.6924 | -21.9013
13.5026 | -34.6929 | -21.9097
13.5087 | -34.6933 | -21.9179
13.5148 | -34.6937 | -21.9264
13.5209 | -34.6942 | -21.9348
13.527 | -34.6946 | -21.9432
13.5331 | -34.6951 | -21.9515
13.5393 | -34.6955 | -21.9595
13.5454 | -34.696 | -21.9679
13.5515 | -34.6965 | -21.9765
13.5576 | -34.6969 | -21.9848
13.5637 | -34.6974 | -21.9932
13.5698 | -34.6978 | -22.0015
13.5759 | -34.6983 | -22.0097
13.582 | -34.6987 | -22.018
13.5881 | -34.6992 | -22.0263
13.5942 | -34.6997 | -22.0348
13.6003 | -34.7001 | -22.0431
13.6064 | -34.7006 | -22.0517
13.6125 | -34.7011 | -22.06
13.6186 | -34.7015 | -22.0682
13.6247 | -34.702 | -22.0766
13.6308 | -34.7025 | -22.085
13.6369 | -34.7029 | -22.0935
13.643 | -34.7033 | -22.102
13.6491 | -34.7038 | -22.1101
13.6552 | -34.7043 | -22.1185
13.6613 | -34.7048 | -22.1268
13.6674 | -34.7052 | -22.1354
13.6735 | -34.7057 | -22.1438
13.6796 | -34.7061 | -22.1521
13.6858 | -34.7066 | -22.1603
13.6919 | -34.707 | -22.1687
13.698 | -34.7075 | -22.1771
13.7041 | -34.708 | -22.1856
13.7102 | -34.7084 | -22.194
13.7163 | -34.7089 | -22.2024
13.7224 | -34.7093 | -22.2109
13.7285 | -34.7098 | -22.2194
13.7346 | -34.7103 | -22.2278
13.7407 | -34.7108 | -22.2362
13.7468 | -34.7112 | -22.2446
13.7529 | -34.7116 | -22.2528
13.759 | -34.7121 | -22.2613
13.7651 | -34.7125 | -22.2697
13.7712 | -34.713 | -22.278
13.7773 | -34.7134 | -22.2865
13.7834 | -34.7139 | -22.2948
13.7895 | -34.7144 | -22.3031
13.7956 | -34.7148 | -22.3115
13.8017 | -34.7153 | -22.3198
13.8078 | -34.7158 | -22.3282
13.8139 | -34.7162 | -22.3365
13.82 | -34.7167 | -22.345
13.8262 | -34.7171 | -22.3535
13.8323 | -34.7176 | -22.3618
13.8384 | -34.718 | -22.3703
13.8445 | -34.7185 | -22.3787
13.8506 | -34.719 | -22.3871
13.8567 | -34.7194 | -22.3956
13.8628 | -34.7199 | -22.404
13.8689 | -34.7204 | -22.4125
13.875 | -34.7208 | -22.4208
13.8811 | -34.7212 | -22.4291
13.8872 | -34.7217 | -22.4375
13.8933 | -34.7221 | -22.4457
13.8994 | -34.7226 | -22.4542
13.9055 | -34.723 | -22.4627
13.9116 | -34.7235 | -22.471
13.9177 | -34.7239 | -22.4794
13.9238 | -34.7244 | -22.4877
13.9299 | -34.7249 | -22.4959
13.936 | -34.7253 | -22.5043
13.9421 | -34.7258 | -22.5127
13.9482 | -34.7262 | -22.521
13.9543 | -34.7266 | -22.5295
13.9604 | -34.7271 | -22.5378
13.9665 | -34.7276 | -22.546
13.9727 | -34.728 | -22.5543
13.9788 | -34.7285 | -22.5626
13.9849 | -34.729 | -22.571
13.991 | -34.7294 | -22.5793
13.9971 | -34.7299 | -22.5874
14.0032 | -34.7303 | -22.5957
14.0093 | -34.7308 | -22.6041
14.0154 | -34.7312 | -22.6126
14.0215 | -34.7317 | -22.6207
14.0276 | -34.7322 | -22.629
14.0337 | -34.7326 | -22.6374
14.0398 | -34.7331 | -22.6456
14.0459 | -34.7335 | -22.6538
14.052 | -34.734 | -22.6622
14.0581 | -34.7344 | -22.6707
14.0642 | -34.7349 | -22.679
14.0703 | -34.7353 | -22.6872
14.0764 | -34.7358 | -22.6954
14.0825 | -34.7362 | -22.7038
14.0886 | -34.7367 | -22.7121
14.0947 | -34.7372 | -22.7204
14.1008 | -34.7376 | -22.7287
14.1069 | -34.738 | -22.737
14.1131 | -34.7385 | -22.7455
14.1192 | -34.739 | -22.7537
14.1253 | -34.7394 | -22.7619
14.1314 | -34.7399 | -22.7703
14.1375 | -34.7403 | -22.7785
14.1436 | -34.7408 | -22.7869
14.1497 | -34.7413 | -22.7952
14.1558 | -34.7417 | -22.8036
14.1619 | -34.7422 | -22.8117
14.168 | -34.7426 | -22.8201
14.1741 | -34.7431 | -22.8282
14.1802 | -34.7435 | -22.8364
14.1863 | -34.744 | -22.8447
14.1924 | -34.7445 | -22.853
14.1985 | -34.7449 | -22.8612
14.2046 | -34.7454 | -22.8696
14.2107 | -34.7458 | -22.8777
14.2168 | -34.7463 | -22.8861
14.2229 | -34.7467 | -22.8944
14.229 | -34.7472 | -22.9027
14.2351 | -34.7476 | -22.911
14.2412 | -34.7481 | -22.9192
14.2473 | -34.7486 | -22.9274
14.2534 | -34.749 | -22.9356
14.2596 | -34.7495 | -22.9438
14.2657 | -34.7499 | -22.9521
14.2718 | -34.7504 | -22.9604
14.2779 | -34.7508 | -22.9687
14.284 | -34.7513 | -22.9769
14.2901 | -34.7518 | -22.9852
14.2962 | -34.7522 | -22.9934
14.3023 | -34.7527 | -23.0018
14.3084 | -34.7531 | -23.01
14.3145 | -34.7536 | -23.0182
14.3206 | -34.754 | -23.0267
14.3267 | -34.7545 | -23.035
14.3328 | -34.755 | -23.0432
14.3389 | -34.7554 | -23.0513
14.345 | -34.7559 | -23.0596
14.3511 | -34.7563 | -23.0679
14.3572 | -34.7568 | -23.0762
14.3633 | -34.7573 | -23.0846
14.3694 | -34.7578 | -23.0927
14.3755 | -34.7582 | -23.1011
14.3816 | -34.7587 | -23.1095
14.3877 | -34.7592 | -23.1177
14.3938 | -34.7596 | -23.1261
14.4 | -34.7601 | -23.1343
14.4061 | -34.7605 | -23.1426
14.4122 | -34.761 | -23.1509
14.4183 | -34.7614 | -23.159
14.4244 | -34.7619 | -23.1673
14.4305 | -34.7624 | -23.1755
14.4366 | -34.7628 | -23.1837
14.4427 | -34.7633 | -23.1921
14.4488 | -34.7638 | -23.2005
14.4549 | -34.7642 | -23.2086
14.461 | -34.7647 | -23.2167
14.4671 | -34.7651 | -23.225
14.4732 | -34.7656 | -23.2333
14.4793 | -34.7661 | -23.2416
14.4854 | -34.7665 | -23.2498
14.4915 | -34.767 | -23.258
14.4976 | -34.7674 | -23.2661
14.5037 | -34.7679 | -23.2743
14.5098 | -34.7684 | -23.2826
14.5159 | -34.7688 | -23.2908
14.522 | -34.7692 | -23.299
14.5281 | -34.7697 | -23.307
14.5342 | -34.7702 | -23.3153
14.5403 | -34.7706 | -23.3235
14.5465 | -34.7711 | -23.3317
14.5526 | -34.7715 | -23.3401
14.5587 | -34.772 | -23.3485
14.5648 | -34.7725 | -23.3567
14.5709 | -34.773 | -23.3648
14.577 | -34.7734 | -23.3731
14.5831 | -34.7738 | -23.3814
14.5892 | -34.7743 | -23.3896
14.5953 | -34.7747 | -23.3977
14.6014 | -34.7752 | -23.4059
14.6075 | -34.7756 | -23.4142
14.6136 | -34.7761 | -23.4225
14.6197 | -34.7765 | -23.4307
14.6258 | -34.777 | -23.4389
14.6319 | -34.7775 | -23.4472
14.638 | -34.7779 | -23.4553
14.6441 | -34.7784 | -23.4636
14.6502 | -34.7788 | -23.4719
14.6563 | -34.7793 | -23.4801
14.6624 | -34.7798 | -23.4884
14.6685 | -34.7803 | -23.4966
14.6746 | -34.7807 | -23.5048
14.6807 | -34.7812 | -23.5131
14.6869 | -34.7816 | -23.5213
14.693 | -34.7821 | -23.5294
14.6991 | -34.7825 | -23.5376
14.7052 | -34.783 | -23.5456
14.7113 | -34.7835 | -23.5538
14.7174 | -34.7839 | -23.5621
14.7235 | -34.7844 | -23.5702
14.7296 | -34.7849 | -23.5783
14.7357 | -34.7853 | -23.5864
14.7418 | -34.7858 | -23.5944
14.7479 | -34.7863 | -23.6025
14.754 | -34.7867 | -23.6107
14.7601 | -34.7872 | -23.6189
14.7662 | -34.7876 | -23.627
14.7723 | -34.7881 | -23.6349
14.7784 | -34.7885 | -23.6432
14.7845 | -34.789 | -23.6513
14.7906 | -34.7895 | -23.6596
14.7967 | -34.7899 | -23.6677
14.8028 | -34.7904 | -23.6758
14.8089 | -34.7908 | -23.6838
14.815 | -34.7913 | -23.6921
14.8211 | -34.7918 | -23.7003
14.8272 | -34.7922 | -23.7085
14.8334 | -34.7927 | -23.7166
14.8395 | -34.7931 | -23.7248
14.8456 | -34.7936 | -23.7328
14.8517 | -34.7941 | -23.7411
14.8578 | -34.7945 | -23.7494
14.8639 | -34.795 | -23.7577
14.87 | -34.7955 | -23.7658
14.8761 | -34.796 | -23.774
14.8822 | -34.7964 | -23.7822
14.8883 | -34.7969 | -23.7905
14.8944 | -34.7973 | -23.7985
14.9005 | -34.7978 | -23.8067
14.9066 | -34.7982 | -23.815
14.9127 | -34.7987 | -23.8231
14.9188 | -34.7991 | -23.8314
14.9249 | -34.7996 | -23.8396
14.931 | -34.8001 | -23.8479
14.9371 | -34.8006 | -23.8559
14.9432 | -34.801 | -23.8639
14.9493 | -34.8015 | -23.8722
14.9554 | -34.8019 | -23.8803
14.9615 | -34.8024 | -23.8885
14.9676 | -34.8028 | -23.8967
14.9738 | -34.8033 | -23.9049
14.9799 | -34.8038 | -23.913
14.986 | -34.8042 | -23.921
14.9921 | -34.8046 | -23.929
14.9982 | -34.8051 | -23.9371
15.0043 | -34.8055 | -23.9452
15.0104 | -34.806 | -23.9534
15.0165 | -34.8065 | -23.9616
15.0226 | -34.8069 | -23.9693
15.0287 | -34.8074 | -23.9774
15.0348 | -34.8079 | -23.9855
15.0409 | -34.8083 | -23.9939
15.047 | -34.8088 | -24.002
15.0531 | -34.8093 | -24.0101
15.0592 | -34.8097 | -24.0181
15.0653 | -34.8102 | -24.0263
15.0714 | -34.8106 | -24.0342
15.0775 | -34.8111 | -24.0422
15.0836 | -34.8116 | -24.0502
15.0897 | -34.8121 | -24.0584
15.0958 | -34.8125 | -24.0663
15.1019 | -34.8129 | -24.0742
15.108 | -34.8134 | -24.0823
15.1141 | -34.8139 | -24.0903
15.1203 | -34.8144 | -24.0984
15.1264 | -34.8149 | -24.1066
15.1325 | -34.8153 | -24.1149
15.1386 | -34.8158 | -24.1231
15.1447 | -34.8162 | -24.1314
15.1508 | -34.8167 | -24.1397
15.1569 | -34.8171 | -24.1477
15.163 | -34.8176 | -24.1557
15.1691 | -34.8181 | -24.1638
15.1752 | -34.8185 | -24.1718
15.1813 | -34.819 | -24.1799
15.1874 | -34.8195 | -24.188
15.1935 | -34.8199 | -24.196
15.1996 | -34.8203 | -24.2041
15.2057 | -34.8208 | -24.212
15.2118 | -34.8212 | -24.22
15.2179 | -34.8217 | -24.2281
15.224 | -34.8222 | -24.2361
15.2301 | -34.8227 | -24.244
15.2362 | -34.8231 | -24.2522
15.2423 | -34.8236 | -24.2602
15.2484 | -34.824 | -24.2684
15.2545 | -34.8245 | -24.2765
15.2607 | -34.825 | -24.2844
15.2668 | -34.8255 | -24.2924
15.2729 | -34.826 | -24.3002
15.279 | -34.8265 | -24.3085
15.2851 | -34.8269 | -24.3162
15.2912 | -34.8274 | -24.3242
15.2973 | -34.8279 | -24.3321
15.3034 | -34.8284 | -24.3399
15.3095 | -34.8288 | -24.3481
15.3156 | -34.8293 | -24.356
15.3217 | -34.8298 | -24.3641
15.3278 | -34.8302 | -24.3721
15.3339 | -34.8307 | -24.3799
15.34 | -34.8312 | -24.3878
15.3461 | -34.8317 | -24.3956
15.3522 | -34.8321 | -24.4035
15.3583 | -34.8326 | -24.4116
15.3644 | -34.833 | -24.4196
15.3705 | -34.8336 | -24.4272
15.3766 | -34.834 | -24.4353
15.3827 | -34.8345 | -24.4432
15.3888 | -34.8349 | -24.4511
15.3949 | -34.8354 | -24.459
15.401 | -34.8359 | -24.4672
15.4072 | -34.8364 | -24.4751
15.4133 | -34.8368 | -24.4833
15.4194 | -34.8373 | -24.4913
15.4255 | -34.8377 | -24.4993
15.4316 | -34.8382 | -24.5075
15.4377 | -34.8387 | -24.5155
15.4438 | -34.8393 | -24.5232
15.4499 | -34.8397 | -24.5312
15.456 | -34.8402 | -24.5389
15.4621 | -34.8407 | -24.547
15.4682 | -34.8412 | -24.555
15.4743 | -34.8416 | -24.5626
15.4804 | -34.8421 | -24.5705
15.4865 | -34.8425 | -24.5786
15.4926 | -34.843 | -24.5865
15.4987 | -34.8435 | -24.5944
15.5048 | -34.844 | -24.6023
15.5109 | -34.8444 | -24.6101
15.517 | -34.8449 | -24.618
15.5231 | -34.8454 | -24.626
15.5292 | -34.8458 | -24.6338
15.5353 | -34.8463 | -24.6418
15.5414 | -34.8468 | -24.6496
15.5476 | -34.8473 | -24.6568
15.5537 | -34.8477 | -24.6647
15.5598 | -34.8482 | -24.6726
15.5659 | -34.8487 | -24.6804
15.572 | -34.8491 | -24.6882
15.5781 | -34.8496 | -24.696
15.5842 | -34.8501 | -24.7037
15.5903 | -34.8506 | -24.7113
15.5964 | -34.8511 | -24.7192
15.6025 | -34.8515 | -24.7269
15.6086 | -34.852 | -24.735
15.6147 | -34.8525 | -24.7429
15.6208 | -34.853 | -24.7506
15.6269 | -34.8534 | -24.7584
15.633 | -34.8539 | -24.7665
15.6391 | -34.8543 | -24.7743
15.6452 | -34.8548 | -24.7822
15.6513 | -34.8554 | -24.7902
15.6574 | -34.8558 | -24.7981
15.6635 | -34.8563 | -24.806
15.6696 | -34.8567 | -24.8141
15.6757 | -34.8572 | -24.822
15.6818 | -34.8577 | -24.8299
15.688 | -34.8582 | -24.8377
15.6941 | -34.8587 | -24.8456
15.7002 | -34.8591 | -24.8534
15.7063 | -34.8596 | -24.8611
15.7124 | -34.86 | -24.8691
15.7185 | -34.8605 | -24.8771
15.7246 | -34.861 | -24.885
15.7307 | -34.8615 | -24.8928
15.7368 | -34.8619 | -24.9005
15.7429 | -34.8624 | -24.9083
15.749 | -34.8629 | -24.9162
15.7551 | -34.8634 | -24.924
15.7612 | -34.8639 | -24.9319
15.7673 | -34.8643 | -24.9396
15.7734 | -34.8648 | -24.9473
15.7795 | -34.8653 | -24.9551
15.7856 | -34.8658 | -24.9626
15.7917 | -34.8663 | -24.9698
15.7978 | -34.8668 | -24.9778
15.8039 | -34.8672 | -24.9857
15.81 | -34.8677 | -24.9937
15.8161 | -34.8682 | -25.0015
15.8222 | -34.8687 | -25.0093
15.8283 | -34.8691 | -25.0171
15.8345 | -34.8696 | -25.025
15.8406 | -34.8701 | -25.0327
15.8467 | -34.8706 | -25.0407
15.8528 | -34.8711 | -25.0485
15.8589 | -34.8716 | -25.0564
15.865 | -34.8721 | -25.0637
15.8711 | -34.8726 | -25.0717
15.8772 | -34.873 | -25.0793
15.8833 | -34.8735 | -25.0873
15.8894 | -34.8739 | -25.0951
15.8955 | -34.8744 | -25.1029
15.9016 | -34.8749 | -25.1104
15.9077 | -34.8754 | -25.1181
15.9138 | -34.8758 | -25.1259
15.9199 | -34.8763 | -25.1337
15.926 | -34.8768 | -25.1413
15.9321 | -34.8773 | -25.1489
15.9382 | -34.8778 | -25.1565
15.9443 | -34.8782 | -25.1644
15.9504 | -34.8787 | -25.1722
15.9565 | -34.8792 | -25.1799
15.9626 | -34.8797 | -25.1878
15.9687 | -34.8802 | -25.1953
15.9749 | -34.8806 | -25.203
15.981 | -34.8811 | -25.2107
15.9871 | -34.8816 | -25.2186
15.9932 | -34.882 | -25.2262
15.9993 | -34.8825 | -25.2339
16.0054 | -34.883 | -25.2414
16.0115 | -34.8835 | -25.2493
16.0176 | -34.884 | -25.2572
16.0237 | -34.8844 | -25.2651
16.0298 | -34.8849 | -25.2727
16.0359 | -34.8854 | -25.2804
16.042 | -34.8859 | -25.2881
16.0481 | -34.8863 | -25.2958
16.0542 | -34.8868 | -25.3035
16.0603 | -34.8873 | -25.3112
16.0664 | -34.8878 | -25.3188
16.0725 | -34.8882 | -25.3261
16.0786 | -34.8887 | -25.3337
16.0847 | -34.8892 | -25.3415
16.0908 | -34.8897 | -25.3493
16.0969 | -34.8901 | -25.3571
16.103 | -34.8906 | -25.3648
16.1091 | -34.8911 | -25.3721
16.1152 | -34.8916 | -25.3796
16.1214 | -34.8921 | -25.3874
16.1275 | -34.8926 | -25.395
16.1336 | -34.8931 | -25.4026
16.1397 | -34.8935 | -25.4102
16.1458 | -34.894 | -25.4176
16.1519 | -34.8945 | -25.4254
16.158 | -34.895 | -25.4331
16.1641 | -34.8954 | -25.4409
16.1702 | -34.8959 | -25.4486
16.1763 | -34.8964 | -25.4562
16.1824 | -34.8968 | -25.4637
16.1885 | -34.8973 | -25.4714
16.1946 | -34.8978 | -25.4792
16.2007 | -34.8983 | -25.4868
16.2068 | -34.8988 | -25.4945
16.2129 | -34.8993 | -25.5022
16.219 | -34.8998 | -25.5096
16.2251 | -34.9003 | -25.5172
16.2312 | -34.9008 | -25.5247
16.2373 | -34.9012 | -25.5321
16.2434 | -34.9017 | -25.5397
16.2495 | -34.9022 | -25.547
16.2556 | -34.9027 | -25.5547
16.2618 | -34.9032 | -25.5622
16.2679 | -34.9037 | -25.5698
16.274 | -34.9042 | -25.5774
16.2801 | -34.9046 | -25.5849
16.2862 | -34.9051 | -25.5924
16.2923 | -34.9057 | -25.6
16.2984 | -34.9061 | -25.6076
16.3045 | -34.9066 | -25.6153
16.3106 | -34.9071 | -25.6228
16.3167 | -34.9076 | -25.6305
16.3228 | -34.908 | -25.6379
16.3289 | -34.9085 | -25.6455
16.335 | -34.909 | -25.653
16.3411 | -34.9095 | -25.6607
16.3472 | -34.91 | -25.6685
16.3533 | -34.9105 | -25.6761
16.3594 | -34.911 | -25.6833
16.3655 | -34.9114 | -25.691
16.3716 | -34.9119 | -25.6985
16.3777 | -34.9123 | -25.706
16.3838 | -34.9128 | -25.7136
16.3899 | -34.9133 | -25.7212
16.396 | -34.9138 | -25.7289
16.4021 | -34.9143 | -25.7364
16.4083 | -34.9147 | -25.7439
16.4144 | -34.9152 | -25.7515
16.4205 | -34.9157 | -25.7591
16.4266 | -34.9162 | -25.7665
16.4327 | -34.9166 | -25.774
16.4388 | -34.9171 | -25.7816
16.4449 | -34.9176 | -25.7891
16.451 | -34.9181 | -25.7967
16.4571 | -34.9186 | -25.8044
16.4632 | -34.919 | -25.8117
16.4693 | -34.9195 | -25.8192
16.4754 | -34.92 | -25.827
16.4815 | -34.9205 | -25.8346
16.4876 | -34.9209 | -25.8421
16.4937 | -34.9214 | -25.8495
16.4998 | -34.9219 | -25.857
16.5059 | -34.9224 | -25.8646
16.512 | -34.9229 | -25.8722
16.5181 | -34.9234 | -25.8799
16.5242 | -34.9239 | -25.8876
16.5303 | -34.9243 | -25.895
16.5364 | -34.9248 | -25.9025
16.5425 | -34.9253 | -25.91
16.5487 | -34.9258 | -25.9175
16.5548 | -34.9263 | -25.9251
16.5609 | -34.9268 | -25.9327
16.567 | -34.9273 | -25.9398
16.5731 | -34.9278 | -25.9474
16.5792 | -34.9283 | -25.955
16.5853 | -34.9287 | -25.9624
16.5914 | -34.9292 | -25.9699
16.5975 | -34.9297 | -25.9775
16.6036 | -34.9302 | -25.9849
16.6097 | -34.9307 | -25.9925
16.6158 | -34.9312 | -26
16.6219 | -34.9317 | -26.0076
16.628 | -34.9321 | -26.0151
16.6341 | -34.9326 | -26.0226
16.6402 | -34.9331 | -26.03
16.6463 | -34.9336 | -26.0375
16.6524 | -34.934 | -26.0449
16.6585 | -34.9345 | -26.0525
16.6646 | -34.935 | -26.06
16.6707 | -34.9355 | -26.0673
16.6768 | -34.936 | -26.0748
16.6829 | -34.9365 | -26.0823
16.689 | -34.937 | -26.0899
16.6952 | -34.9375 | -26.0973
16.7013 | -34.9379 | -26.1048
16.7074 | -34.9385 | -26.112
16.7135 | -34.9389 | -26.1194
16.7196 | -34.9394 | -26.1269
16.7257 | -34.9399 | -26.1343
16.7318 | -34.9404 | -26.1416
16.7379 | -34.9409 | -26.1491
16.744 | -34.9414 | -26.1565
16.7501 | -34.9419 | -26.164
16.7562 | -34.9424 | -26.1715
16.7623 | -34.9428 | -26.1789
16.7684 | -34.9433 | -26.1865
16.7745 | -34.9438 | -26.194
16.7806 | -34.9443 | -26.2011
16.7867 | -34.9448 | -26.2085
16.7928 | -34.9453 | -26.2159
16.7989 | -34.9458 | -26.2233
16.805 | -34.9462 | -26.2308
16.8111 | -34.9467 | -26.2381
16.8172 | -34.9472 | -26.2457
16.8233 | -34.9477 | -26.2532
16.8294 | -34.9482 | -26.2606
16.8356 | -34.9486 | -26.2681
16.8417 | -34.9491 | -26.2755
16.8478 | -34.9497 | -26.2827
16.8539 | -34.9501 | -26.2901
16.86 | -34.9506 | -26.2976
16.8661 | -34.9511 | -26.305
16.8722 | -34.9516 | -26.3124
16.8783 | -34.9521 | -26.3199
16.8844 | -34.9526 | -26.3272
16.8905 | -34.9531 | -26.3348
16.8966 | -34.9535 | -26.3422
16.9027 | -34.954 | -26.3496
16.9088 | -34.9545 | -26.3571
16.9149 | -34.955 | -26.3645
16.921 | -34.9555 | -26.3718
16.9271 | -34.956 | -26.3792
16.9332 | -34.9565 | -26.3867
16.9393 | -34.957 | -26.3942
16.9454 | -34.9575 | -26.4016
16.9515 | -34.9579 | -26.409
16.9576 | -34.9584 | -26.4162
16.9637 | -34.9589 | -26.4236
16.9698 | -34.9594 | -26.4309
16.9759 | -34.9599 | -26.4383
16.9821 | -34.9604 | -26.4458
16.9882 | -34.9609 | -26.4529
16.9943 | -34.9614 | -26.4603
17.0004 | -34.9618 | -26.4676
17.0065 | -34.9623 | -26.4751
17.0126 | -34.9628 | -26.4825
17.0187 | -34.9633 | -26.49
17.0248 | -34.9638 | -26.4973
17.0309 | -34.9643 | -26.5048
17.037 | -34.9648 | -26.5121
17.0431 | -34.9653 | -26.5194
17.0492 | -34.9658 | -26.5269
17.0553 | -34.9662 | -26.5343
17.0614 | -34.9667 | -26.5415
17.0675 | -34.9672 | -26.5489
17.0736 | -34.9677 | -26.5565
17.0797 | -34.9682 | -26.5638
17.0858 | -34.9687 | -26.5714
17.0919 | -34.9692 | -26.5788
17.098 | -34.9696 | -26.586
17.1041 | -34.9701 | -26.5933
17.1102 | -34.9706 | -26.6006
17.1163 | -34.9711 | -26.6079
17.1225 | -34.9716 | -26.6153
17.1286 | -34.9721 | -26.6225
17.1347 | -34.9726 | -26.6299
17.1408 | -34.9731 | -26.6374
17.1469 | -34.9736 | -26.6447
17.153 | -34.9741 | -26.6521
17.1591 | -34.9746 | -26.6593
17.1652 | -34.9751 | -26.6664
17.1713 | -34.9755 | -26.6736
17.1774 | -34.976 | -26.681
17.1835 | -34.9765 | -26.6882
17.1896 | -34.977 | -26.6956
17.1957 | -34.9775 | -26.7028
17.2018 | -34.978 | -26.7097
17.2079 | -34.9785 | -26.7171
17.214 | -34.979 | -26.7245
17.2201 | -34.9795 | -26.7318
17.2262 | -34.98 | -26.7392
17.2323 | -34.9805 | -26.7466
17.2384 | -34.981 | -26.7539
17.2445 | -34.9815 | -26.761
17.2506 | -34.982 | -26.7683
17.2567 | -34.9824 | -26.7758
17.2628 | -34.9829 | -26.7831
17.269 | -34.9834 | -26.7904
17.2751 | -34.9839 | -26.7976
17.2812 | -34.9844 | -26.805
17.2873 | -34.9849 | -26.8124
17.2934 | -34.9854 | -26.8198
17.2995 | -34.9859 | -26.8273
17.3056 | -34.9864 | -26.8345
17.3117 | -34.9869 | -26.8419
17.3178 | -34.9874 | -26.8492
17.3239 | -34.9879 | -26.8565
17.33 | -34.9883 | -26.8638
17.3361 | -34.9888 | -26.8711
17.3422 | -34.9893 | -26.8785
17.3483 | -34.9898 | -26.8859
17.3544 | -34.9903 | -26.8933
17.3605 | -34.9908 | -26.9006
17.3666 | -34.9913 | -26.9079
17.3727 | -34.9917 | -26.9152
17.3788 | -34.9922 | -26.9223
17.3849 | -34.9927 | -26.9296
17.391 | -34.9932 | -26.9369
17.3971 | -34.9937 | -26.9442
17.4032 | -34.9942 | -26.9515
17.4094 | -34.9947 | -26.9586
17.4155 | -34.9952 | -26.9659
17.4216 | -34.9956 | -26.9731
17.4277 | -34.9961 | -26.9804
17.4338 | -34.9966 | -26.9876
17.4399 | -34.9971 | -26.9951
17.446 | -34.9976 | -27.0022
17.4521 | -34.9981 | -27.0096
17.4582 | -34.9986 | -27.0168
17.4643 | -34.9991 | -27.0241
17.4704 | -34.9996 | -27.0313
17.4765 | -35 | -27.0387
17.4826 | -35.0005 | -27.0459
17.4887 | -35.001 | -27.0533
17.4948 | -35.0015 | -27.0606
17.5009 | -35.002 | -27.0678
17.507 | -35.0025 | -27.0752
17.5131 | -35.003 | -27.0824
17.5192 | -35.0035 | -27.0896
17.5253 | -35.004 | -27.097
17.5314 | -35.0044 | -27.1042
17.5375 | -35.0049 | -27.1115
17.5436 | -35.0054 | -27.1189
17.5497 | -35.0059 | -27.1263
17.5559 | -35.0064 | -27.1334
17.562 | -35.0069 | -27.1407
17.5681 | -35.0074 | -27.148
17.5742 | -35.0079 | -27.1553
17.5803 | -35.0084 | -27.1626
17.5864 | -35.0089 | -27.1698
17.5925 | -35.0094 | -27.1771
17.5986 | -35.0099 | -27.1843
17.6047 | -35.0104 | -27.1917
17.6108 | -35.0109 | -27.1989
17.6169 | -35.0114 | -27.2062
17.623 | -35.0119 | -27.2134
17.6291 | -35.0124 | -27.2206
17.6352 | -35.0128 | -27.2279
17.6413 | -35.0133 | -27.2352
17.6474 | -35.0138 | -27.2425
17.6535 | -35.0143 | -27.2496
17.6596 | -35.0148 | -27.2568
17.6657 | -35.0153 | -27.2641
17.6718 | -35.0158 | -27.2714
17.6779 | -35.0163 | -27.2786
17.684 | -35.0168 | -27.286
17.6901 | -35.0173 | -27.2932
17.6963 | -35.0178 | -27.3004
17.7024 | -35.0183 | -27.3077
17.7085 | -35.0188 | -27.3148
17.7146 | -35.0193 | -27.3221
17.7207 | -35.0198 | -27.3294
17.7268 | -35.0203 | -27.3366
17.7329 | -35.0207 | -27.344
17.739 | -35.0212 | -27.3512
17.7451 | -35.0217 | -27.3584
17.7512 | -35.0222 | -27.3657
17.7573 | -35.0227 | -27.3731
17.7634 | -35.0232 | -27.3801
17.7695 | -35.0237 | -27.3874
17.7756 | -35.0242 | -27.3947
17.7817 | -35.0247 | -27.402
17.7878 | -35.0252 | -27.4093
17.7939 | -35.0257 | -27.4166
17.8 | -35.0262 | -27.4236
17.8061 | -35.0267 | -27.431
17.8122 | -35.0272 | -27.4382
17.8183 | -35.0277 | -27.4455
17.8244 | -35.0282 | -27.4527
17.8305 | -35.0287 | -27.46
17.8366 | -35.0292 | -27.4672
17.8428 | -35.0297 | -27.4744
17.8489 | -35.0302 | -27.4815
17.855 | -35.0306 | -27.4887
17.8611 | -35.0311 | -27.4959
17.8672 | -35.0316 | -27.5029
17.8733 | -35.0321 | -27.5102
17.8794 | -35.0326 | -27.5175
17.8855 | -35.0331 | -27.5247
17.8916 | -35.0336 | -27.532
17.8977 | -35.0341 | -27.5392
17.9038 | -35.0346 | -27.5463
17.9099 | -35.0351 | -27.5534
17.916 | -35.0356 | -27.5607
17.9221 | -35.0361 | -27.568
17.9282 | -35.0366 | -27.5753
17.9343 | -35.0371 | -27.5826
17.9404 | -35.0375 | -27.5897
17.9465 | -35.038 | -27.5969
17.9526 | -35.0385 | -27.6042
17.9587 | -35.039 | -27.6114
17.9648 | -35.0395 | -27.6184
17.9709 | -35.04 | -27.6257
17.977 | -35.0405 | -27.6328
17.9832 | -35.041 | -27.6401
17.9893 | -35.0414 | -27.6474
17.9954 | -35.0419 | -27.6546
18.0015 | -35.0424 | -27.6619
18.0076 | -35.0429 | -27.6689
18.0137 | -35.0434 | -27.6762
18.0198 | -35.0439 | -27.6834
18.0259 | -35.0444 | -27.6907
18.032 | -35.0448 | -27.6978
18.0381 | -35.0453 | -27.7051
18.0442 | -35.0458 | -27.7123
18.0503 | -35.0463 | -27.7195
18.0564 | -35.0468 | -27.7268
18.0625 | -35.0473 | -27.7341
18.0686 | -35.0478 | -27.7414
18.0747 | -35.0483 | -27.7487
18.0808 | -35.0487 | -27.7558
18.0869 | -35.0493 | -27.763
18.093 | -35.0498 | -27.7703
18.0991 | -35.0502 | -27.7775
18.1052 | -35.0507 | -27.7848
18.1113 | -35.0512 | -27.7919
18.1174 | -35.0517 | -27.7988
18.1236 | -35.0522 | -27.806
18.1297 | -35.0527 | -27.8133
18.1358 | -35.0532 | -27.8205
18.1419 | -35.0537 | -27.8277
18.148 | -35.0542 | -27.8348
18.1541 | -35.0547 | -27.8419
18.1602 | -35.0552 | -27.8491
18.1663 | -35.0557 | -27.8562
18.1724 | -35.0561 | -27.8634
18.1785 | -35.0566 | -27.8706
18.1846 | -35.0571 | -27.8778
18.1907 | -35.0576 | -27.885
18.1968 | -35.0581 | -27.8922
18.2029 | -35.0586 | -27.8995
18.209 | -35.059 | -27.9068
18.2151 | -35.0595 | -27.9141
18.2212 | -35.06 | -27.9211
18.2273 | -35.0605 | -27.9284
18.2334 | -35.061 | -27.9356
18.2395 | -35.0615 | -27.9429
18.2456 | -35.062 | -27.9501
18.2517 | -35.0624 | -27.9573
18.2578 | -35.0629 | -27.9646
18.2639 | -35.0634 | -27.9718
18.2701 | -35.0639 | -27.9791
18.2762 | -35.0644 | -27.9862
18.2823 | -35.0649 | -27.9933
18.2884 | -35.0654 | -28.0004
18.2945 | -35.0658 | -28.0073
18.3006 | -35.0663 | -28.0147
18.3067 | -35.0668 | -28.0219
18.3128 | -35.0673 | -28.0292
18.3189 | -35.0678 | -28.0364
18.325 | -35.0683 | -28.0434
18.3311 | -35.0688 | -28.0505
18.3372 | -35.0693 | -28.0576
18.3433 | -35.0697 | -28.0649
18.3494 | -35.0702 | -28.0721
18.3555 | -35.0707 | -28.0792
18.3616 | -35.0712 | -28.0861
18.3677 | -35.0717 | -28.0932
18.3738 | -35.0722 | -28.1002
18.3799 | -35.0727 | -28.1073
18.386 | -35.0732 | -28.1146
18.3921 | -35.0737 | -28.1218
18.3982 | -35.0741 | -28.1287
18.4043 | -35.0746 | -28.1359
18.4105 | -35.0751 | -28.1431
18.4166 | -35.0756 | -28.1502
18.4227 | -35.0761 | -28.1574
18.4288 | -35.0765 | -28.1645
18.4349 | -35.077 | -28.1716
18.441 | -35.0775 | -28.1787
18.4471 | -35.078 | -28.1859
18.4532 | -35.0784 | -28.1932
18.4593 | -35.0789 | -28.2003
18.4654 | -35.0794 | -28.2074
18.4715 | -35.0799 | -28.2146
18.4776 | -35.0804 | -28.2219
18.4837 | -35.0809 | -28.2291
18.4898 | -35.0814 | -28.2363
18.4959 | -35.0819 | -28.2437
18.502 | -35.0823 | -28.2508
18.5081 | -35.0828 | -28.258
18.5142 | -35.0833 | -28.2652
18.5203 | -35.0838 | -28.2723
18.5264 | -35.0843 | -28.2794
18.5325 | -35.0848 | -28.2866
18.5386 | -35.0852 | -28.2937
18.5447 | -35.0857 | -28.3008
18.5508 | -35.0862 | -28.3078
18.557 | -35.0867 | -28.3152
18.5631 | -35.0871 | -28.3224
18.5692 | -35.0876 | -28.3295
18.5753 | -35.0881 | -28.3365
18.5814 | -35.0886 | -28.3437
18.5875 | -35.0891 | -28.3507
18.5936 | -35.0895 | -28.3579
18.5997 | -35.09 | -28.3651
18.6058 | -35.0905 | -28.3722
18.6119 | -35.091 | -28.3793
18.618 | -35.0914 | -28.3864
18.6241 | -35.0919 | -28.3935
18.6302 | -35.0924 | -28.4008
18.6363 | -35.0929 | -28.4079
18.6424 | -35.0934 | -28.415
18.6485 | -35.0938 | -28.4221
18.6546 | -35.0943 | -28.4292
18.6607 | -35.0948 | -28.4364
18.6668 | -35.0953 | -28.4436
18.6729 | -35.0957 | -28.4507
18.679 | -35.0962 | -28.4577
18.6851 | -35.0967 | -28.4647
18.6912 | -35.0972 | -28.4719
18.6974 | -35.0977 | -28.479
18.7035 | -35.0981 | -28.4862
18.7096 | -35.0986 | -28.4933
18.7157 | -35.0991 | -28.5004
18.7218 | -35.0996 | -28.5075
18.7279 | -35.1001 | -28.5146
18.734 | -35.1005 | -28.5217
18.7401 | -35.101 | -28.5287
18.7462 | -35.1015 | -28.5356
18.7523 | -35.102 | -28.5428
18.7584 | -35.1025 | -28.5499
18.7645 | -35.1029 | -28.557
18.7706 | -35.1034 | -28.5641
18.7767 | -35.1039 | -28.5712
18.7828 | -35.1043 | -28.5781
18.7889 | -35.1048 | -28.5854
18.795 | -35.1053 | -28.5927
18.8011 | -35.1058 | -28.5998
18.8072 | -35.1062 | -28.607
18.8133 | -35.1067 | -28.614
18.8194 | -35.1071 | -28.6209
18.8255 | -35.1076 | -28.6279
18.8316 | -35.1081 | -28.635
18.8377 | -35.1086 | -28.6421
18.8439 | -35.109 | -28.6492
18.85 | -35.1095 | -28.6564
18.8561 | -35.11 | -28.6634
18.8622 | -35.1104 | -28.6705
18.8683 | -35.1109 | -28.6775
18.8744 | -35.1114 | -28.6845
18.8805 | -35.1119 | -28.6918
18.8866 | -35.1124 | -28.6989
18.8927 | -35.1129 | -28.7058
18.8988 | -35.1133 | -28.7128
18.9049 | -35.1138 | -28.7199
18.911 | -35.1143 | -28.7269
18.9171 | -35.1148 | -28.7338
18.9232 | -35.1153 | -28.7407
18.9293 | -35.1158 | -28.7478
18.9354 | -35.1163 | -28.755
18.9415 | -35.1167 | -28.7621
18.9476 | -35.1172 | -28.7691
18.9537 | -35.1177 | -28.7762
18.9598 | -35.1181 | -28.783
18.9659 | -35.1186 | -28.7901
18.972 | -35.1191 | -28.7972
18.9781 | -35.1195 | -28.8043
18.9843 | -35.12 | -28.8115
18.9904 | -35.1204 | -28.8185
18.9965 | -35.1209 | -28.8254
19.0026 | -35.1214 | -28.8325
19.0087 | -35.1218 | -28.8396
19.0148 | -35.1223 | -28.8465
19.0209 | -35.1228 | -28.8537
19.027 | -35.1233 | -28.8608
19.0331 | -35.1237 | -28.8677
19.0392 | -35.1242 | -28.8749
19.0453 | -35.1247 | -28.882
19.0514 | -35.1251 | -28.8891
19.0575 | -35.1256 | -28.8963
19.0636 | -35.1261 | -28.9029
19.0697 | -35.1266 | -28.9099
19.0758 | -35.127 | -28.9169
19.0819 | -35.1275 | -28.9241
19.088 | -35.128 | -28.9311
19.0941 | -35.1284 | -28.9384
19.1002 | -35.1289 | -28.9453
19.1063 | -35.1293 | -28.9523
19.1124 | -35.1298 | -28.9593
19.1185 | -35.1303 | -28.9662
19.1246 | -35.1308 | -28.9733
19.1308 | -35.1312 | -28.9803
19.1369 | -35.1317 | -28.9872
19.143 | -35.1322 | -28.9942
19.1491 | -35.1326 | -29.0013
19.1552 | -35.1331 | -29.0083
19.1613 | -35.1335 | -29.0154
19.1674 | -35.134 | -29.0224
19.1735 | -35.1345 | -29.0293
19.1796 | -35.135 | -29.0364
19.1857 | -35.1354 | -29.0434
19.1918 | -35.1359 | -29.0504
19.1979 | -35.1364 | -29.0576
19.204 | -35.1369 | -29.0644
19.2101 | -35.1373 | -29.0714
19.2162 | -35.1378 | -29.0784
19.2223 | -35.1383 | -29.0854
19.2284 | -35.1388 | -29.0925
19.2345 | -35.1392 | -29.0996
19.2406 | -35.1397 | -29.1064
19.2467 | -35.1402 | -29.1135
19.2528 | -35.1406 | -29.1205
19.2589 | -35.1411 | -29.1273
19.265 | -35.1415 | -29.1344
19.2712 | -35.142 | -29.1414
19.2773 | -35.1424 | -29.1482
19.2834 | -35.1429 | -29.1551
19.2895 | -35.1434 | -29.162
19.2956 | -35.1439 | -29.1691
19.3017 | -35.1443 | -29.176
19.3078 | -35.1448 | -29.1831
19.3139 | -35.1452 | -29.1899
19.32 | -35.1457 | -29.1969
19.3261 | -35.1462 | -29.2038
19.3322 | -35.1466 | -29.2107
19.3383 | -35.1471 | -29.2176
19.3444 | -35.1476 | -29.2244
19.3505 | -35.148 | -29.2314
19.3566 | -35.1484 | -29.2384
19.3627 | -35.1489 | -29.2454
19.3688 | -35.1494 | -29.2523
19.3749 | -35.1499 | -29.2592
19.381 | -35.1503 | -29.2659
19.3871 | -35.1508 | -29.2728
19.3932 | -35.1512 | -29.2799
19.3993 | -35.1517 | -29.2869
19.4054 | -35.1521 | -29.2939
19.4115 | -35.1526 | -29.301
19.4177 | -35.153 | -29.3077
19.4238 | -35.1535 | -29.3145
19.4299 | -35.154 | -29.3213
19.436 | -35.1544 | -29.3283
19.4421 | -35.1549 | -29.3352
19.4482 | -35.1553 | -29.3422
19.4543 | -35.1558 | -29.3489
19.4604 | -35.1562 | -29.3558
19.4665 | -35.1567 | -29.3627
19.4726 | -35.1571 | -29.3695
19.4787 | -35.1576 | -29.3763
19.4848 | -35.158 | -29.3832
19.4909 | -35.1585 | -29.39
19.497 | -35.159 | -29.3967
19.5031 | -35.1594 | -29.4035
19.5092 | -35.1599 | -29.4103
19.5153 | -35.1603 | -29.4172
19.5214 | -35.1608 | -29.4239
19.5275 | -35.1612 | -29.4308
19.5336 | -35.1616 | -29.4377
19.5397 | -35.1621 | -29.4447
19.5458 | -35.1626 | -29.4516
19.5519 | -35.163 | -29.4584
19.5581 | -35.1634 | -29.465
19.5642 | -35.1639 | -29.4719
19.5703 | -35.1644 | -29.4789
19.5764 | -35.1648 | -29.4859
19.5825 | -35.1653 | -29.4927
19.5886 | -35.1657 | -29.4995
19.5947 | -35.1662 | -29.5063
19.6008 | -35.1666 | -29.5131
19.6069 | -35.1671 | -29.52
19.613 | -35.1675 | -29.527
19.6191 | -35.168 | -29.5338
19.6252 | -35.1684 | -29.5407
19.6313 | -35.1689 | -29.5473
19.6374 | -35.1693 | -29.5541
19.6435 | -35.1698 | -29.5609
19.6496 | -35.1702 | -29.5678
19.6557 | -35.1707 | -29.5748
19.6618 | -35.1711 | -29.5814
19.6679 | -35.1716 | -29.5883
19.674 | -35.172 | -29.5951
19.6801 | -35.1725 | -29.602
19.6862 | -35.1729 | -29.6088
19.6923 | -35.1734 | -29.6156
19.6984 | -35.1738 | -29.6222
19.7046 | -35.1743 | -29.6291
19.7107 | -35.1748 | -29.6359
19.7168 | -35.1752 | -29.6429
19.7229 | -35.1757 | -29.6497
19.729 | -35.1761 | -29.6565
19.7351 | -35.1766 | -29.6631
19.7412 | -35.177 | -29.6701
19.7473 | -35.1775 | -29.6769
19.7534 | -35.1779 | -29.6837
19.7595 | -35.1784 | -29.6905
19.7656 | -35.1788 | -29.6974
19.7717 | -35.1793 | -29.7039
19.7778 | -35.1797 | -29.7107
19.7839 | -35.1801 | -29.7176
19.79 | -35.1806 | -29.7245
19.7961 | -35.181 | -29.7314
19.8022 | -35.1814 | -29.7379
19.8083 | -35.1819 | -29.7447
19.8144 | -35.1823 | -29.7515
19.8205 | -35.1828 | -29.7583
19.8266 | -35.1832 | -29.7651
19.8327 | -35.1837 | -29.7719
19.8388 | -35.1841 | -29.7784
19.845 | -35.1846 | -29.7853
19.8511 | -35.185 | -29.7922
19.8572 | -35.1854 | -29.7991
19.8633 | -35.1859 | -29.8059
19.8694 | -35.1863 | -29.8127
19.8755 | -35.1867 | -29.8193
19.8816 | -35.1872 | -29.8263
19.8877 | -35.1876 | -29.8332
19.8938 | -35.188 | -29.8401
19.8999 | -35.1885 | -29.8469
19.906 | -35.1889 | -29.8536
19.9121 | -35.1893 | -29.8602
19.9182 | -35.1898 | -29.867
19.9243 | -35.1902 | -29.8738
19.9304 | -35.1906 | -29.8805
19.9365 | -35.1911 | -29.8873
19.9426 | -35.1915 | -29.8939
19.9487 | -35.1919 | -29.9007
19.9548 | -35.1924 | -29.9076
19.9609 | -35.1929 | -29.9145
19.967 | -35.1933 | -29.9211
19.9731 | -35.1938 | -29.9278
19.9792 | -35.1942 | -29.9345
19.9853 | -35.1946 | -29.9412
19.9915 | -35.1951 | -29.948
19.9976 | -35.1955 | -29.9548
20.0037 | -35.196 | -29.9617
20.0098 | -35.1964 | -29.9684
20.0159 | -35.1968 | -29.975
20.022 | -35.1973 | -29.9817
20.0281 | -35.1977 | -29.9885
20.0342 | -35.1982 | -29.9952
20.0403 | -35.1986 | -30.002
20.0464 | -35.1991 | -30.0087
20.0525 | -35.1995 | -30.0153
20.0586 | -35.2 | -30.0219
20.0647 | -35.2004 | -30.0288
20.0708 | -35.2008 | -30.0355
20.0769 | -35.2013 | -30.0421
20.083 | -35.2017 | -30.0486
20.0891 | -35.2021 | -30.0553
20.0952 | -35.2026 | -30.062
20.1013 | -35.203 | -30.0687
20.1074 | -35.2035 | -30.0753
20.1135 | -35.2039 | -30.082
20.1196 | -35.2043 | -30.0883
20.1257 | -35.2047 | -30.0951
20.1319 | -35.2052 | -30.1019
20.138 | -35.2056 | -30.1086
20.1441 | -35.2061 | -30.1154
20.1502 | -35.2065 | -30.1221
20.1563 | -35.207 | -30.1285
20.1624 | -35.2074 | -30.1353
20.1685 | -35.2078 | -30.1421
20.1746 | -35.2083 | -30.1489
20.1807 | -35.2087 | -30.1556
20.1868 | -35.2091 | -30.1624
20.1929 | -35.2095 | -30.1688
20.199 | -35.21 | -30.1755
20.2051 | -35.2104 | -30.1821
20.2112 | -35.2109 | -30.1887
20.2173 | -35.2113 | -30.1953
20.2234 | -35.2117 | -30.2021
20.2295 | -35.2122 | -30.2085
20.2356 | -35.2126 | -30.2153
20.2417 | -35.213 | -30.222
20.2478 | -35.2135 | -30.2288
20.2539 | -35.2139 | -30.2357
20.26 | -35.2143 | -30.242
20.2661 | -35.2148 | -30.2487
20.2723 | -35.2152 | -30.2553
20.2784 | -35.2157 | -30.262
20.2845 | -35.2161 | -30.2686
20.2906 | -35.2165 | -30.2753
20.2967 | -35.2169 | -30.2819
20.3028 | -35.2174 | -30.2885
20.3089 | -35.2178 | -30.2951
20.315 | -35.2183 | -30.3017
20.3211 | -35.2187 | -30.3084
20.3272 | -35.2191 | -30.3151
20.3333 | -35.2195 | -30.3216
20.3394 | -35.22 | -30.3282
20.3455 | -35.2204 | -30.3349
20.3516 | -35.2209 | -30.3414
20.3577 | -35.2213 | -30.348
20.3638 | -35.2217 | -30.3546
20.3699 | -35.2221 | -30.3609
20.376 | -35.2226 | -30.3676
20.3821 | -35.223 | -30.3743
20.3882 | -35.2234 | -30.381
20.3943 | -35.2239 | -30.3876
20.4004 | -35.2243 | -30.394
20.4065 | -35.2248 | -30.4007
20.4126 | -35.2252 | -30.4073
20.4188 | -35.2256 | -30.414
20.4249 | -35.2261 | -30.4205
20.431 | -35.2265 | -30.427
20.4371 | -35.2269 | -30.4334
20.4432 | -35.2274 | -30.4399
20.4493 | -35.2278 | -30.4467
20.4554 | -35.2282 | -30.4532
20.4615 | -35.2286 | -30.4599
20.4676 | -35.2291 | -30.4665
20.4737 | -35.2295 | -30.4729
20.4798 | -35.2299 | -30.4795
20.4859 | -35.2304 | -30.4862
20.492 | -35.2308 | -30.4927
20.4981 | -35.2312 | -30.4993
20.5042 | -35.2317 | -30.5059
20.5103 | -35.2321 | -30.5121
20.5164 | -35.2325 | -30.5186
20.5225 | -35.233 | -30.5252
20.5286 | -35.2334 | -30.5318
20.5347 | -35.2338 | -30.5383
20.5408 | -35.2343 | -30.5447
20.5469 | -35.2347 | -30.5514
20.553 | -35.2351 | -30.5579
20.5592 | -35.2356 | -30.5645
20.5653 | -35.236 | -30.571
20.5714 | -35.2364 | -30.5777
20.5775 | -35.2369 | -30.584
20.5836 | -35.2373 | -30.5905
20.5897 | -35.2377 | -30.5971
20.5958 | -35.2382 | -30.6038
20.6019 | -35.2386 | -30.6105
20.608 | -35.239 | -30.6171
20.6141 | -35.2395 | -30.6235
20.6202 | -35.2399 | -30.63
20.6263 | -35.2403 | -30.6365
20.6324 | -35.2408 | -30.6431
20.6385 | -35.2412 | -30.6498
20.6446 | -35.2416 | -30.6563
20.6507 | -35.242 | -30.6628
20.6568 | -35.2425 | -30.6694
20.6629 | -35.2429 | -30.6758
20.669 | -35.2433 | -30.6825
20.6751 | -35.2438 | -30.6888
20.6812 | -35.2442 | -30.695
20.6873 | -35.2446 | -30.7014
20.6934 | -35.245 | -30.7079
20.6995 | -35.2455 | -30.7143
20.7057 | -35.2459 | -30.721
20.7118 | -35.2463 | -30.7274
20.7179 | -35.2467 | -30.7335
20.724 | -35.2472 | -30.7401
20.7301 | -35.2476 | -30.7465
20.7362 | -35.248 | -30.7529
20.7423 | -35.2485 | -30.7594
20.7484 | -35.2489 | -30.766
20.7545 | -35.2493 | -30.7725
20.7606 | -35.2497 | -30.7789
20.7667 | -35.2502 | -30.7853
20.7728 | -35.2506 | -30.7918
20.7789 | -35.2511 | -30.7983
20.785 | -35.2515 | -30.8049
20.7911 | -35.2519 | -30.8114
20.7972 | -35.2523 | -30.8181
20.8033 | -35.2527 | -30.8246
20.8094 | -35.2532 | -30.831
20.8155 | -35.2536 | -30.8375
20.8216 | -35.254 | -30.8438
20.8277 | -35.2545 | -30.85
20.8338 | -35.2549 | -30.8563
20.8399 | -35.2553 | -30.8629
20.8461 | -35.2557 | -30.8692
20.8522 | -35.2562 | -30.8758
20.8583 | -35.2566 | -30.882
20.8644 | -35.257 | -30.8883
20.8705 | -35.2575 | -30.8949
20.8766 | -35.2579 | -30.9014
20.8827 | -35.2583 | -30.9077
20.8888 | -35.2587 | -30.9141
20.8949 | -35.2591 | -30.9203
20.901 | -35.2595 | -30.9268
20.9071 | -35.26 | -30.9332
20.9132 | -35.2604 | -30.9398
20.9193 | -35.2608 | -30.9463
20.9254 | -35.2613 | -30.9527
20.9315 | -35.2617 | -30.959
20.9376 | -35.2621 | -30.9653
20.9437 | -35.2625 | -30.9718
20.9498 | -35.2629 | -30.9783
20.9559 | -35.2634 | -30.9848
20.962 | -35.2638 | -30.9913
20.9681 | -35.2642 | -30.9975
20.9742 | -35.2647 | -31.004
20.9803 | -35.2651 | -31.0105
20.9864 | -35.2655 | -31.0169
20.9926 | -35.2659 | -31.0235
20.9987 | -35.2663 | -31.0298
21.0048 | -35.2668 | -31.0362
21.0109 | -35.2672 | -31.0428
21.017 | -35.2676 | -31.0491
21.0231 | -35.2681 | -31.0555
21.0292 | -35.2685 | -31.0617
21.0353 | -35.2689 | -31.0677
21.0414 | -35.2693 | -31.0741
21.0475 | -35.2697 | -31.0808
21.0536 | -35.2701 | -31.0871
21.0597 | -35.2706 | -31.0937
21.0658 | -35.271 | -31.1002
21.0719 | -35.2714 | -31.1065
21.078 | -35.2719 | -31.113
21.0841 | -35.2723 | -31.1193
21.0902 | -35.2727 | -31.1255
21.0963 | -35.2731 | -31.1323
21.1024 | -35.2736 | -31.1388
21.1085 | -35.274 | -31.1449
21.1146 | -35.2745 | -31.1512
21.1207 | -35.2749 | -31.1576
21.1268 | -35.2753 | -31.1641
21.133 | -35.2757 | -31.1703
21.1391 | -35.2761 | -31.1764
21.1452 | -35.2765 | -31.1827
21.1513 | -35.277 | -31.1888
21.1574 | -35.2774 | -31.1953
21.1635 | -35.2778 | -31.2018
21.1696 | -35.2783 | -31.2082
21.1757 | -35.2787 | -31.2145
21.1818 | -35.2791 | -31.2209
21.1879 | -35.2795 | -31.2271
21.194 | -35.2799 | -31.2336
21.2001 | -35.2804 | -31.2398
21.2062 | -35.2808 | -31.2461
21.2123 | -35.2812 | -31.2521
21.2184 | -35.2816 | -31.2583
21.2245 | -35.282 | -31.2646
21.2306 | -35.2824 | -31.271
21.2367 | -35.2829 | -31.2773
21.2428 | -35.2833 | -31.2833
21.2489 | -35.2837 | -31.2895
21.255 | -35.2841 | -31.2959
21.2611 | -35.2845 | -31.3023
21.2672 | -35.285 | -31.3086
21.2733 | -35.2854 | -31.3149
21.2795 | -35.2858 | -31.3207
21.2856 | -35.2862 | -31.327
21.2917 | -35.2867 | -31.3332
21.2978 | -35.2871 | -31.3399
21.3039 | -35.2876 | -31.3464
21.31 | -35.288 | -31.3529
21.3161 | -35.2884 | -31.3591
21.3222 | -35.2888 | -31.3653
21.3283 | -35.2892 | -31.3716
21.3344 | -35.2896 | -31.3778
21.3405 | -35.29 | -31.3842
21.3466 | -35.2905 | -31.3907
21.3527 | -35.2909 | -31.3967
21.3588 | -35.2914 | -31.4032
21.3649 | -35.2918 | -31.4096
21.371 | -35.2922 | -31.4157
21.3771 | -35.2926 | -31.4221
21.3832 | -35.293 | -31.4285
21.3893 | -35.2935 | -31.4344
21.3954 | -35.2939 | -31.4406
21.4015 | -35.2943 | -31.447
21.4076 | -35.2948 | -31.4532
21.4137 | -35.2952 | -31.4595
21.4199 | -35.2957 | -31.4659
21.426 | -35.2961 | -31.4717
21.4321 | -35.2965 | -31.4782
21.4382 | -35.2969 | -31.4844
21.4443 | -35.2973 | -31.4903
21.4504 | -35.2978 | -31.4968
21.4565 | -35.2982 | -31.5028
21.4626 | -35.2986 | -31.5093
21.4687 | -35.299 | -31.5156
21.4748 | -35.2995 | -31.5219
21.4809 | -35.2999 | -31.528
21.487 | -35.3003 | -31.534
21.4931 | -35.3007 | -31.5401
21.4992 | -35.3012 | -31.5464
21.5053 | -35.3016 | -31.5525
21.5114 | -35.302 | -31.5588
21.5175 | -35.3024 | -31.5649
21.5236 | -35.3029 | -31.571
21.5297 | -35.3033 | -31.577
21.5358 | -35.3037 | -31.5832
21.5419 | -35.3042 | -31.5896
21.548 | -35.3046 | -31.5958
21.5541 | -35.305 | -31.602
21.5602 | -35.3054 | -31.6084
21.5664 | -35.3058 | -31.6145
21.5725 | -35.3062 | -31.6208
21.5786 | -35.3066 | -31.6271
21.5847 | -35.307 | -31.6334
21.5908 | -35.3074 | -31.6394
21.5969 | -35.3078 | -31.6454
21.603 | -35.3083 | -31.6517
21.6091 | -35.3087 | -31.6581
21.6152 | -35.3091 | -31.6646
21.6213 | -35.3095 | -31.6709
21.6274 | -35.31 | -31.6774
21.6335 | -35.3104 | -31.6838
21.6396 | -35.3108 | -31.6899
21.6457 | -35.3112 | -31.6962
21.6518 | -35.3116 | -31.7024
21.6579 | -35.3121 | -31.7088
21.664 | -35.3125 | -31.7151
21.6701 | -35.3129 | -31.7209
21.6762 | -35.3133 | -31.727
21.6823 | -35.3137 | -31.7336
21.6884 | -35.3142 | -31.74
21.6945 | -35.3146 | -31.7462
21.7006 | -35.315 | -31.7526
21.7068 | -35.3154 | -31.7587
21.7129 | -35.3158 | -31.765
21.719 | -35.3163 | -31.7714
21.7251 | -35.3167 | -31.7777
21.7312 | -35.3171 | -31.7842
21.7373 | -35.3175 | -31.7901
21.7434 | -35.318 | -31.7964
21.7495 | -35.3184 | -31.8029
21.7556 | -35.3188 | -31.809
21.7617 | -35.3192 | -31.8153
21.7678 | -35.3196 | -31.8214
21.7739 | -35.3201 | -31.8275
21.78 | -35.3205 | -31.8338
21.7861 | -35.3209 | -31.84
21.7922 | -35.3213 | -31.8462
21.7983 | -35.3217 | -31.8525
21.8044 | -35.3222 | -31.8585
21.8105 | -35.3226 | -31.8646
21.8166 | -35.323 | -31.871
21.8227 | -35.3234 | -31.8771
21.8288 | -35.3238 | -31.8834
21.8349 | -35.3242 | -31.8897
21.841 | -35.3246 | -31.896
21.8471 | -35.3251 | -31.9019
21.8533 | -35.3255 | -31.9083
21.8594 | -35.3259 | -31.9146
21.8655 | -35.3263 | -31.9208
21.8716 | -35.3267 | -31.9272
21.8777 | -35.3272 | -31.9333
21.8838 | -35.3277 | -31.9393
21.8899 | -35.3281 | -31.9452
21.896 | -35.3285 | -31.9514
21.9021 | -35.329 | -31.9575
21.9082 | -35.3294 | -31.9636
21.9143 | -35.3298 | -31.9696
21.9204 | -35.3302 | -31.9757
21.9265 | -35.3307 | -31.9816
21.9326 | -35.3311 | -31.9878
21.9387 | -35.3315 | -31.9939
21.9448 | -35.3319 | -32.0001
21.9509 | -35.3323 | -32.0061
21.957 | -35.3327 | -32.0124
21.9631 | -35.3331 | -32.0187
21.9692 | -35.3335 | -32.0247
21.9753 | -35.334 | -32.0309
21.9814 | -35.3344 | -32.0372
21.9875 | -35.3349 | -32.0432
21.9937 | -35.3353 | -32.0495
21.9998 | -35.3357 | -32.0559
22.0059 | -35.3361 | -32.062
22.012 | -35.3366 | -32.0682
22.0181 | -35.337 | -32.0741
22.0242 | -35.3374 | -32.0802
22.0303 | -35.3378 | -32.0867
22.0364 | -35.3383 | -32.0928
22.0425 | -35.3387 | -32.0988
22.0486 | -35.3391 | -32.105
22.0547 | -35.3395 | -32.1108
22.0608 | -35.3399 | -32.1168
22.0669 | -35.3404 | -32.123
22.073 | -35.3408 | -32.1293
22.0791 | -35.3412 | -32.1355
22.0852 | -35.3416 | -32.142
22.0913 | -35.342 | -32.148
22.0974 | -35.3424 | -32.1542
22.1035 | -35.3428 | -32.1604
22.1096 | -35.3432 | -32.1667
22.1157 | -35.3436 | -32.1729
22.1218 | -35.344 | -32.1789
22.1279 | -35.3445 | -32.1848
22.134 | -35.3449 | -32.1909
22.1402 | -35.3453 | -32.197
22.1463 | -35.3457 | -32.2031
22.1524 | -35.3461 | -32.2093
22.1585 | -35.3465 | -32.2155
22.1646 | -35.3469 | -32.2215
22.1707 | -35.3473 | -32.2278
22.1768 | -35.3477 | -32.2339
22.1829 | -35.3481 | -32.2401
22.189 | -35.3485 | -32.246
22.1951 | -35.3489 | -32.2519
22.2012 | -35.3494 | -32.2582
22.2073 | -35.3498 | -32.2643
22.2134 | -35.3502 | -32.2704
22.2195 | -35.3507 | -32.2767
22.2256 | -35.3511 | -32.2828
22.2317 | -35.3515 | -32.2888
22.2378 | -35.3519 | -32.2946
22.2439 | -35.3523 | -32.3005
22.25 | -35.3527 | -32.3068
22.2561 | -35.3531 | -32.3127
22.2622 | -35.3535 | -32.3186
22.2683 | -35.354 | -32.3246
22.2744 | -35.3544 | -32.3307
22.2806 | -35.3548 | -32.3369
22.2867 | -35.3552 | -32.343
22.2928 | -35.3556 | -32.3491
22.2989 | -35.3561 | -32.3552
22.305 | -35.3565 | -32.3612
22.3111 | -35.3569 | -32.3672
22.3172 | -35.3573 | -32.3734
22.3233 | -35.3578 | -32.3797
22.3294 | -35.3582 | -32.3856
22.3355 | -35.3586 | -32.3915
22.3416 | -35.359 | -32.3975
22.3477 | -35.3594 | -32.4035
22.3538 | -35.3598 | -32.4095
22.3599 | -35.3602 | -32.4156
22.366 | -35.3606 | -32.4215
22.3721 | -35.3611 | -32.4274
22.3782 | -35.3615 | -32.4335
22.3843 | -35.3619 | -32.4397
22.3904 | -35.3623 | -32.4455
22.3965 | -35.3627 | -32.4515
22.4026 | -35.3631 | -32.4576
22.4087 | -35.3635 | -32.4635
22.4148 | -35.3639 | -32.4697
22.4209 | -35.3644 | -32.4759
22.4271 | -35.3647 | -32.4822
22.4332 | -35.3652 | -32.4883
22.4393 | -35.3656 | -32.4945
22.4454 | -35.366 | -32.5003
22.4515 | -35.3665 | -32.5064
22.4576 | -35.3669 | -32.5124
22.4637 | -35.3673 | -32.5184
22.4698 | -35.3677 | -32.5243
22.4759 | -35.3681 | -32.5301
22.482 | -35.3685 | -32.5363
22.4881 | -35.369 | -32.5424
22.4942 | -35.3694 | -32.5483
22.5003 | -35.3698 | -32.5543
22.5064 | -35.3702 | -32.5603
22.5125 | -35.3706 | -32.5664
22.5186 | -35.371 | -32.5723
22.5247 | -35.3714 | -32.5783
22.5308 | -35.3718 | -32.5844
22.5369 | -35.3722 | -32.5906
22.543 | -35.3726 | -32.5967
22.5491 | -35.3731 | -32.6026
22.5552 | -35.3735 | -32.6085
22.5613 | -35.3739 | -32.6148
22.5675 | -35.3744 | -32.6207
22.5736 | -35.3748 | -32.6266
22.5797 | -35.3752 | -32.6327
22.5858 | -35.3756 | -32.6384
22.5919 | -35.3761 | -32.6445
22.598 | -35.3765 | -32.6506
22.6041 | -35.3769 | -32.6568
22.6102 | -35.3773 | -32.6629
22.6163 | -35.3777 | -32.6688
22.6224 | -35.3782 | -32.6747
22.6285 | -35.3786 | -32.6808
22.6346 | -35.379 | -32.6867
22.6407 | -35.3794 | -32.6926
22.6468 | -35.3798 | -32.6987
22.6529 | -35.3803 | -32.7044
22.659 | -35.3807 | -32.7102
22.6651 | -35.3811 | -32.7163
22.6712 | -35.3815 | -32.7222
22.6773 | -35.3819 | -32.7282
22.6834 | -35.3824 | -32.7342
22.6895 | -35.3828 | -32.7399
22.6956 | -35.3832 | -32.746
22.7017 | -35.3836 | -32.7519
22.7079 | -35.384 | -32.758
22.714 | -35.3844 | -32.7638
22.7201 | -35.3849 | -32.7699
22.7262 | -35.3853 | -32.7759
22.7323 | -35.3857 | -32.782
22.7384 | -35.3862 | -32.7878
22.7445 | -35.3865 | -32.7939
22.7506 | -35.387 | -32.7997
22.7567 | -35.3874 | -32.8056
22.7628 | -35.3879 | -32.8114
22.7689 | -35.3883 | -32.8172
22.775 | -35.3887 | -32.8231
22.7811 | -35.3891 | -32.8291
22.7872 | -35.3895 | -32.8352
22.7933 | -35.3899 | -32.8411
22.7994 | -35.3904 | -32.8471
22.8055 | -35.3908 | -32.8531
22.8116 | -35.3912 | -32.8592
22.8177 | -35.3917 | -32.8651
22.8238 | -35.3921 | -32.8713
22.8299 | -35.3925 | -32.8772
22.836 | -35.3929 | -32.8832
22.8421 | -35.3933 | -32.8893
22.8482 | -35.3937 | -32.8952
22.8544 | -35.3941 | -32.9012
22.8605 | -35.3945 | -32.9072
22.8666 | -35.3949 | -32.9131
22.8727 | -35.3953 | -32.9192
22.8788 | -35.3957 | -32.925
22.8849 | -35.3961 | -32.931
22.891 | -35.3966 | -32.937
22.8971 | -35.397 | -32.943
22.9032 | -35.3974 | -32.9488
22.9093 | -35.3978 | -32.9548
22.9154 | -35.3982 | -32.9609
22.9215 | -35.3986 | -32.9669
22.9276 | -35.399 | -32.9729
22.9337 | -35.3994 | -32.9786
22.9398 | -35.3998 | -32.9846
22.9459 | -35.4003 | -32.9907
22.952 | -35.4007 | -32.9965
22.9581 | -35.4011 | -33.0026
22.9642 | -35.4015 | -33.0085
22.9703 | -35.4019 | -33.0141
22.9764 | -35.4023 | -33.0202
22.9825 | -35.4027 | -33.0262
22.9886 | -35.4032 | -33.0323
22.9948 | -35.4036 | -33.0383
23.0009 | -35.404 | -33.0442
23.007 | -35.4044 | -33.0499
23.0131 | -35.4048 | -33.0557
23.0192 | -35.4053 | -33.0617
23.0253 | -35.4057 | -33.0676
23.0314 | -35.406 | -33.0738
23.0375 | -35.4065 | -33.0797
23.0436 | -35.4069 | -33.0856
23.0497 | -35.4073 | -33.0916
23.0558 | -35.4077 | -33.0975
23.0619 | -35.4081 | -33.1035
23.068 | -35.4085 | -33.1095
23.0741 | -35.4089 | -33.1151
23.0802 | -35.4093 | -33.1212
23.0863 | -35.4097 | -33.1271
23.0924 | -35.4101 | -33.133
23.0985 | -35.4105 | -33.139
23.1046 | -35.4109 | -33.1449
23.1107 | -35.4113 | -33.1506
23.1168 | -35.4118 | -33.1567
23.1229 | -35.4122 | -33.1626
23.129 | -35.4126 | -33.1685
23.1351 | -35.413 | -33.1743
23.1413 | -35.4134 | -33.1803
23.1474 | -35.4138 | -33.186
23.1535 | -35.4142 | -33.1921
23.1596 | -35.4146 | -33.198
23.1657 | -35.415 | -33.2037
23.1718 | -35.4154 | -33.2097
23.1779 | -35.4158 | -33.2156
23.184 | -35.4163 | -33.2213
23.1901 | -35.4167 | -33.2272
23.1962 | -35.4171 | -33.2331
23.2023 | -35.4175 | -33.239
23.2084 | -35.4179 | -33.2449
23.2145 | -35.4183 | -33.2504
23.2206 | -35.4187 | -33.2562
23.2267 | -35.4191 | -33.262
23.2328 | -35.4195 | -33.2679
23.2389 | -35.4199 | -33.2739
23.245 | -35.4203 | -33.2798
23.2511 | -35.4207 | -33.2854
23.2572 | -35.4211 | -33.2913
23.2633 | -35.4216 | -33.2973
23.2694 | -35.422 | -33.3034
23.2755 | -35.4224 | -33.3093
23.2817 | -35.4228 | -33.3152
23.2878 | -35.4232 | -33.3209
23.2939 | -35.4236 | -33.3269
23.3 | -35.424 | -33.3328
23.3061 | -35.4244 | -33.3387
23.3122 | -35.4249 | -33.3446
23.3183 | -35.4253 | -33.3506
23.3244 | -35.4257 | -33.3564
23.3305 | -35.4261 | -33.3623
23.3366 | -35.4265 | -33.3681
23.3427 | -35.4269 | -33.3739
23.3488 | -35.4273 | -33.3798
23.3549 | -35.4277 | -33.3857
23.361 | -35.4281 | -33.3915
23.3671 | -35.4286 | -33.3974
23.3732 | -35.429 | -33.4033
23.3793 | -35.4294 | -33.409
23.3854 | -35.4298 | -33.415
23.3915 | -35.4302 | -33.4209
23.3976 | -35.4306 | -33.4268
23.4037 | -35.431 | -33.4327
23.4098 | -35.4314 | -33.4384
23.4159 | -35.4318 | -33.4443
23.422 | -35.4323 | -33.4503
23.4282 | -35.4327 | -33.4559
23.4343 | -35.4331 | -33.4618
23.4404 | -35.4335 | -33.4678
23.4465 | -35.4339 | -33.4736
23.4526 | -35.4343 | -33.4796
23.4587 | -35.4347 | -33.4854
23.4648 | -35.4351 | -33.4911
23.4709 | -35.4355 | -33.497
23.477 | -35.4359 | -33.5027
23.4831 | -35.4363 | -33.5084
23.4892 | -35.4367 | -33.5142
23.4953 | -35.4371 | -33.52
23.5014 | -35.4375 | -33.5257
23.5075 | -35.4379 | -33.5316
23.5136 | -35.4383 | -33.5374
23.5197 | -35.4387 | -33.5432
23.5258 | -35.4391 | -33.5491
23.5319 | -35.4395 | -33.5546
23.538 | -35.4399 | -33.5605
23.5441 | -35.4403 | -33.5664
23.5502 | -35.4407 | -33.5723
23.5563 | -35.4411 | -33.5781
23.5624 | -35.4415 | -33.5841
23.5686 | -35.4419 | -33.5896
23.5747 | -35.4423 | -33.5956
23.5808 | -35.4427 | -33.6013
23.5869 | -35.4431 | -33.6072
23.593 | -35.4435 | -33.613
23.5991 | -35.4439 | -33.6189
23.6052 | -35.4443 | -33.6246
23.6113 | -35.4447 | -33.6305
23.6174 | -35.4451 | -33.6363
23.6235 | -35.4455 | -33.6422
23.6296 | -35.4459 | -33.648
23.6357 | -35.4463 | -33.6536
23.6418 | -35.4467 | -33.6592
23.6479 | -35.4471 | -33.6651
23.654 | -35.4475 | -33.671
23.6601 | -35.4479 | -33.6768
23.6662 | -35.4483 | -33.6827
23.6723 | -35.4487 | -33.6883
23.6784 | -35.4491 | -33.6941
23.6845 | -35.4495 | -33.6998
23.6906 | -35.4499 | -33.7057
23.6967 | -35.4503 | -33.7115
23.7028 | -35.4507 | -33.7173
23.7089 | -35.451 | -33.7228
23.7151 | -35.4514 | -33.7286
23.7212 | -35.4518 | -33.7345
23.7273 | -35.4522 | -33.7403
23.7334 | -35.4526 | -33.7462
23.7395 | -35.453 | -33.7521
23.7456 | -35.4534 | -33.7577
23.7517 | -35.4538 | -33.7636
23.7578 | -35.4542 | -33.7692
23.7639 | -35.4546 | -33.775
23.77 | -35.455 | -33.7808
23.7761 | -35.4554 | -33.7867
23.7822 | -35.4558 | -33.7923
23.7883 | -35.4562 | -33.798
23.7944 | -35.4566 | -33.8037
23.8005 | -35.457 | -33.8096
23.8066 | -35.4574 | -33.8154
23.8127 | -35.4578 | -33.8212
23.8188 | -35.4582 | -33.8271
23.8249 | -35.4586 | -33.8329
23.831 | -35.459 | -33.8388
23.8371 | -35.4594 | -33.8446
23.8432 | -35.4597 | -33.8504
23.8493 | -35.4601 | -33.8559
23.8555 | -35.4605 | -33.8618
23.8616 | -35.4609 | -33.8676
23.8677 | -35.4613 | -33.8734
23.8738 | -35.4617 | -33.8791
23.8799 | -35.4621 | -33.885
23.886 | -35.4625 | -33.8906
23.8921 | -35.4629 | -33.8963
23.8982 | -35.4633 | -33.9022
23.9043 | -35.4637 | -33.908
23.9104 | -35.4641 | -33.9137
23.9165 | -35.4645 | -33.9195
23.9226 | -35.4648 | -33.9251
23.9287 | -35.4652 | -33.9309
23.9348 | -35.4656 | -33.9367
23.9409 | -35.466 | -33.9424
23.947 | -35.4664 | -33.9482
23.9531 | -35.4668 | -33.9537
23.9592 | -35.4671 | -33.9596
23.9653 | -35.4675 | -33.9654
23.9714 | -35.4679 | -33.9712
23.9775 | -35.4683 | -33.977
23.9836 | -35.4687 | -33.9828
23.9897 | -35.4691 | -33.9884
23.9958 | -35.4695 | -33.9942
24.002 | -35.4699 | -33.9998
24.0081 | -35.4703 | -34.0055
24.0142 | -35.4707 | -34.0113
24.0203 | -35.4711 | -34.0171
24.0264 | -35.4714 | -34.0228
24.0325 | -35.4718 | -34.0285
24.0386 | -35.4722 | -34.0344
24.0447 | -35.4726 | -34.0402
24.0508 | -35.473 | -34.0459
24.0569 | -35.4734 | -34.0516
24.063 | -35.4738 | -34.0572
24.0691 | -35.4742 | -34.063
24.0752 | -35.4746 | -34.0687
24.0813 | -35.4749 | -34.0745
24.0874 | -35.4753 | -34.0803
24.0935 | -35.4757 | -34.086
24.0996 | -35.4761 | -34.0916
24.1057 | -35.4765 | -34.0973
24.1118 | -35.4769 | -34.1031
24.1179 | -35.4773 | -34.1089
24.124 | -35.4777 | -34.1146
24.1301 | -35.4781 | -34.1203
24.1362 | -35.4785 | -34.1261
24.1424 | -35.4788 | -34.132
24.1485 | -35.4792 | -34.1377
24.1546 | -35.4796 | -34.1435
24.1607 | -35.48 | -34.1493
24.1668 | -35.4804 | -34.1548
24.1729 | -35.4808 | -34.1606
24.179 | -35.4812 | -34.1663
24.1851 | -35.4815 | -34.172
24.1912 | -35.4819 | -34.1776
24.1973 | -35.4823 | -34.1834
24.2034 | -35.4826 | -34.1889
24.2095 | -35.483 | -34.1948
24.2156 | -35.4834 | -34.2005
24.2217 | -35.4838 | -34.2063
24.2278 | -35.4842 | -34.2121
24.2339 | -35.4846 | -34.2179
24.24 | -35.485 | -34.2235
24.2461 | -35.4854 | -34.2291
24.2522 | -35.4857 | -34.2349
24.2583 | -35.4861 | -34.2405
24.2644 | -35.4865 | -34.2463
24.2705 | -35.4869 | -34.252
24.2766 | -35.4872 | -34.2575
24.2827 | -35.4876 | -34.2634
24.2889 | -35.488 | -34.2692
24.295 | -35.4884 | -34.2748
24.3011 | -35.4887 | -34.2806
24.3072 | -35.4891 | -34.2861
24.3133 | -35.4895 | -34.2919
24.3194 | -35.4899 | -34.2977
24.3255 | -35.4902 | -34.3034
24.3316 | -35.4906 | -34.3091
24.3377 | -35.491 | -34.3146
24.3438 | -35.4914 | -34.3201
24.3499 | -35.4917 | -34.3258
24.356 | -35.4921 | -34.3315
24.3621 | -35.4925 | -34.3372
24.3682 | -35.4929 | -34.3428
24.3743 | -35.4933 | -34.3486
24.3804 | -35.4936 | -34.3541
24.3865 | -35.494 | -34.3598
24.3926 | -35.4944 | -34.3656
24.3987 | -35.4948 | -34.3715
24.4048 | -35.4952 | -34.3772
24.4109 | -35.4955 | -34.3827
24.417 | -35.4959 | -34.3884
24.4231 | -35.4963 | -34.3941
24.4293 | -35.4967 | -34.3999
24.4354 | -35.497 | -34.4057
24.4415 | -35.4974 | -34.4115
24.4476 | -35.4978 | -34.4171
24.4537 | -35.4981 | -34.4228
24.4598 | -35.4985 | -34.4285
24.4659 | -35.4989 | -34.4343
24.472 | -35.4993 | -34.4399
24.4781 | -35.4997 | -34.4456
24.4842 | -35.5 | -34.451
24.4903 | -35.5004 | -34.4567
24.4964 | -35.5008 | -34.4623
24.5025 | -35.5011 | -34.4681
24.5086 | -35.5015 | -34.4736
24.5147 | -35.5019 | -34.4793
24.5208 | -35.5023 | -34.4848
24.5269 | -35.5026 | -34.4905
24.533 | -35.503 | -34.4961
24.5391 | -35.5034 | -34.5018
24.5452 | -35.5038 | -34.5075
24.5513 | -35.5042 | -34.5131
24.5574 | -35.5046 | -34.5187
24.5635 | -35.505 | -34.5245
24.5696 | -35.5053 | -34.5301
24.5758 | -35.5057 | -34.5357
24.5819 | -35.5061 | -34.5412
24.588 | -35.5065 | -34.5467
24.5941 | -35.5069 | -34.5524
24.6002 | -35.5072 | -34.5579
24.6063 | -35.5076 | -34.5635
24.6124 | -35.508 | -34.5691
24.6185 | -35.5084 | -34.5748
24.6246 | -35.5088 | -34.5803
24.6307 | -35.5091 | -34.5859
24.6368 | -35.5095 | -34.5915
24.6429 | -35.5099 | -34.5971
24.649 | -35.5103 | -34.6028
24.6551 | -35.5107 | -34.6084
24.6612 | -35.511 | -34.6139
24.6673 | -35.5114 | -34.6195
24.6734 | -35.5118 | -34.6251
24.6795 | -35.5122 | -34.6308
24.6856 | -35.5125 | -34.6366
24.6917 | -35.5129 | -34.6423
24.6978 | -35.5133 | -34.6476
24.7039 | -35.5137 | -34.6532
24.71 | -35.514 | -34.6589
24.7162 | -35.5144 | -34.6646
24.7223 | -35.5148 | -34.6701
24.7284 | -35.5151 | -34.6756
24.7345 | -35.5155 | -34.6813
24.7406 | -35.5159 | -34.6869
24.7467 | -35.5163 | -34.6926
24.7528 | -35.5166 | -34.6982
24.7589 | -35.517 | -34.7038
24.765 | -35.5174 | -34.7089
24.7711 | -35.5178 | -34.7146
24.7772 | -35.5181 | -34.7202
24.7833 | -35.5185 | -34.7257
24.7894 | -35.5189 | -34.7313
24.7955 | -35.5193 | -34.7368
24.8016 | -35.5196 | -34.7422
24.8077 | -35.52 | -34.7478
24.8138 | -35.5204 | -34.7535
24.8199 | -35.5208 | -34.7592
24.826 | -35.5212 | -34.7648
24.8321 | -35.5215 | -34.7704
24.8382 | -35.5219 | -34.7758
24.8443 | -35.5222 | -34.7813
24.8504 | -35.5226 | -34.7869
24.8565 | -35.523 | -34.7926
24.8627 | -35.5234 | -34.7983
24.8688 | -35.5237 | -34.8038
24.8749 | -35.5241 | -34.8095
24.881 | -35.5245 | -34.8151
24.8871 | -35.5249 | -34.8207
24.8932 | -35.5252 | -34.8264
24.8993 | -35.5256 | -34.8319
24.9054 | -35.5259 | -34.8374
24.9115 | -35.5263 | -34.843
24.9176 | -35.5267 | -34.8486
24.9237 | -35.5271 | -34.8541
24.9298 | -35.5274 | -34.8596
24.9359 | -35.5278 | -34.8653
24.942 | -35.5282 | -34.8706
24.9481 | -35.5285 | -34.8764
24.9542 | -35.5289 | -34.8819
24.9603 | -35.5293 | -34.8875
24.9664 | -35.5296 | -34.8931
24.9725 | -35.53 | -34.8988
24.9786 | -35.5303 | -34.904
24.9847 | -35.5307 | -34.9097
24.9908 | -35.5311 | -34.9153
24.9969 | -35.5314 | -34.9209
25.0031 | -35.5318 | -34.9265
25.0092 | -35.5322 | -34.932
25.0153 | -35.5325 | -34.9374
25.0214 | -35.5329 | -34.9431
25.0275 | -35.5333 | -34.9487
25.0336 | -35.5336 | -34.9542
25.0397 | -35.534 | -34.9599
25.0458 | -35.5344 | -34.9654
25.0519 | -35.5347 | -34.971
25.058 | -35.5351 | -34.9767
25.0641 | -35.5354 | -34.9823
25.0702 | -35.5358 | -34.9878
25.0763 | -35.5362 | -34.9936
25.0824 | -35.5365 | -34.999
25.0885 | -35.5369 | -35.0047
25.0946 | -35.5373 | -35.0103
25.1007 | -35.5377 | -35.0159
25.1068 | -35.538 | -35.0214
25.1129 | -35.5384 | -35.0269
25.119 | -35.5388 | -35.0323
25.1251 | -35.5391 | -35.0378
25.1312 | -35.5395 | -35.0435
25.1373 | -35.5399 | -35.0491
25.1435 | -35.5403 | -35.0548
25.1496 | -35.5406 | -35.06
25.1557 | -35.541 | -35.0655
25.1618 | -35.5413 | -35.0712
25.1679 | -35.5417 | -35.0767
25.174 | -35.542 | -35.0822
25.1801 | -35.5424 | -35.0878
25.1862 | -35.5428 | -35.0932
25.1923 | -35.5431 | -35.0988
25.1984 | -35.5435 | -35.1043
25.2045 | -35.5438 | -35.1099
25.2106 | -35.5442 | -35.1155
25.2167 | -35.5445 | -35.1212
25.2228 | -35.5449 | -35.1267
25.2289 | -35.5452 | -35.1322
25.235 | -35.5456 | -35.1379
25.2411 | -35.546 | -35.1434
25.2472 | -35.5463 | -35.1489
25.2533 | -35.5467 | -35.1544
25.2594 | -35.547 | -35.1595
25.2655 | -35.5474 | -35.1651
25.2716 | -35.5477 | -35.1706
25.2777 | -35.5481 | -35.1762
25.2838 | -35.5484 | -35.1818
25.29 | -35.5488 | -35.1873
25.2961 | -35.5491 | -35.1926
25.3022 | -35.5495 | -35.198
25.3083 | -35.5498 | -35.2036
25.3144 | -35.5502 | -35.2091
25.3205 | -35.5506 | -35.2145
25.3266 | -35.5509 | -35.22
25.3327 | -35.5513 | -35.2254
25.3388 | -35.5516 | -35.231
25.3449 | -35.552 | -35.2364
25.351 | -35.5523 | -35.2419
25.3571 | -35.5527 | -35.2475
25.3632 | -35.553 | -35.2529
25.3693 | -35.5534 | -35.2583
25.3754 | -35.5537 | -35.2638
25.3815 | -35.5541 | -35.2694
25.3876 | -35.5544 | -35.2749
25.3937 | -35.5548 | -35.2804
25.3998 | -35.5551 | -35.2858
25.4059 | -35.5555 | -35.2913
25.412 | -35.5558 | -35.2967
25.4181 | -35.5562 | -35.3022
25.4242 | -35.5565 | -35.3077
25.4304 | -35.5569 | -35.3133
25.4365 | -35.5572 | -35.3187
25.4426 | -35.5576 | -35.3242
25.4487 | -35.5579 | -35.3298
25.4548 | -35.5583 | -35.3352
25.4609 | -35.5587 | -35.3406
25.467 | -35.559 | -35.346
25.4731 | -35.5593 | -35.3515
25.4792 | -35.5597 | -35.3571
25.4853 | -35.5601 | -35.3626
25.4914 | -35.5604 | -35.3681
25.4975 | -35.5608 | -35.3735
25.5036 | -35.5612 | -35.3789
25.5097 | -35.5615 | -35.3843
25.5158 | -35.5619 | -35.3899
25.5219 | -35.5622 | -35.3954
25.528 | -35.5626 | -35.4009
25.5341 | -35.563 | -35.4064
25.5402 | -35.5633 | -35.4117
25.5463 | -35.5637 | -35.4172
25.5524 | -35.564 | -35.4226
25.5585 | -35.5644 | -35.428
25.5646 | -35.5647 | -35.4336
25.5707 | -35.5651 | -35.439
25.5769 | -35.5654 | -35.4444
25.583 | -35.5658 | -35.4498
25.5891 | -35.5661 | -35.4554
25.5952 | -35.5665 | -35.4608
25.6013 | -35.5669 | -35.4663
25.6074 | -35.5672 | -35.4715
25.6135 | -35.5676 | -35.4769
25.6196 | -35.5679 | -35.4824
25.6257 | -35.5683 | -35.4878
25.6318 | -35.5686 | -35.4932
25.6379 | -35.569 | -35.4987
25.644 | -35.5693 | -35.5039
25.6501 | -35.5697 | -35.5093
25.6562 | -35.57 | -35.5146
25.6623 | -35.5704 | -35.5199
25.6684 | -35.5708 | -35.5253
25.6745 | -35.5711 | -35.5307
25.6806 | -35.5715 | -35.536
25.6867 | -35.5718 | -35.5415
25.6928 | -35.5722 | -35.547
25.6989 | -35.5725 | -35.5525
25.705 | -35.5728 | -35.558
25.7111 | -35.5732 | -35.5635
25.7173 | -35.5735 | -35.5688
25.7234 | -35.5739 | -35.5743
25.7295 | -35.5742 | -35.5797
25.7356 | -35.5746 | -35.5854
25.7417 | -35.575 | -35.5909
25.7478 | -35.5753 | -35.5962
25.7539 | -35.5757 | -35.6016
25.76 | -35.576 | -35.607
25.7661 | -35.5764 | -35.6125
25.7722 | -35.5767 | -35.6179
25.7783 | -35.5771 | -35.6234
25.7844 | -35.5774 | -35.6286
25.7905 | -35.5777 | -35.6342
25.7966 | -35.5781 | -35.6396
25.8027 | -35.5784 | -35.6451
25.8088 | -35.5788 | -35.6506
25.8149 | -35.5791 | -35.656
25.821 | -35.5795 | -35.6614
25.8271 | -35.5798 | -35.6669
25.8332 | -35.5802 | -35.6724
25.8393 | -35.5805 | -35.6778
25.8454 | -35.5809 | -35.6833
25.8515 | -35.5812 | -35.6887
25.8576 | -35.5816 | -35.694
25.8638 | -35.5819 | -35.6994
25.8699 | -35.5823 | -35.7047
25.876 | -35.5826 | -35.7102
25.8821 | -35.5829 | -35.7156
25.8882 | -35.5833 | -35.7208
25.8943 | -35.5836 | -35.7263
25.9004 | -35.584 | -35.7317
25.9065 | -35.5843 | -35.7372
25.9126 | -35.5847 | -35.7426
25.9187 | -35.585 | -35.748
25.9248 | -35.5854 | -35.7533
25.9309 | -35.5857 | -35.7587
25.937 | -35.5861 | -35.7641
25.9431 | -35.5865 | -35.7695
25.9492 | -35.5868 | -35.7748
25.9553 | -35.5872 | -35.7802
25.9614 | -35.5875 | -35.7855
25.9675 | -35.5878 | -35.7908
25.9736 | -35.5882 | -35.7963
25.9797 | -35.5885 | -35.8018
25.9858 | -35.5889 | -35.8072
25.9919 | -35.5892 | -35.8127
25.998 | -35.5896 | -35.818
26.0042 | -35.5899 | -35.8234
26.0103 | -35.5902 | -35.8287
26.0164 | -35.5906 | -35.8341
26.0225 | -35.5909 | -35.8394
26.0286 | -35.5913 | -35.8448
26.0347 | -35.5916 | -35.8501
26.0408 | -35.5919 | -35.8555
26.0469 | -35.5922 | -35.8609
26.053 | -35.5926 | -35.8663
26.0591 | -35.593 | -35.8717
26.0652 | -35.5933 | -35.8768
26.0713 | -35.5936 | -35.8821
26.0774 | -35.594 | -35.8875
26.0835 | -35.5943 | -35.8929
26.0896 | -35.5947 | -35.8983
26.0957 | -35.595 | -35.9037
26.1018 | -35.5954 | -35.9089
26.1079 | -35.5957 | -35.9143
26.114 | -35.5961 | -35.9197
26.1201 | -35.5964 | -35.925
26.1262 | -35.5967 | -35.9304
26.1323 | -35.5971 | -35.9358
26.1384 | -35.5974 | -35.941
26.1445 | -35.5978 | -35.9463
26.1507 | -35.5981 | -35.9517
26.1568 | -35.5985 | -35.957
26.1629 | -35.5988 | -35.9625
26.169 | -35.5991 | -35.9679
26.1751 | -35.5995 | -35.973
26.1812 | -35.5998 | -35.9785
26.1873 | -35.6001 | -35.9838
26.1934 | -35.6005 | -35.9891
26.1995 | -35.6008 | -35.9944
26.2056 | -35.6011 | -35.9997
26.2117 | -35.6015 | -36.005
26.2178 | -35.6019 | -36.0103
26.2239 | -35.6022 | -36.0156
26.23 | -35.6025 | -36.021
26.2361 | -35.6028 | -36.0264
26.2422 | -35.6032 | -36.0316
26.2483 | -35.6035 | -36.0369
26.2544 | -35.6039 | -36.0422
26.2605 | -35.6042 | -36.0475
26.2666 | -35.6045 | -36.0529
26.2727 | -35.6049 | -36.0582
26.2788 | -35.6052 | -36.0634
26.2849 | -35.6056 | -36.0687
26.2911 | -35.6059 | -36.0741
26.2972 | -35.6062 | -36.0795
26.3033 | -35.6066 | -36.0849
26.3094 | -35.6069 | -36.0903
26.3155 | -35.6073 | -36.0954
26.3216 | -35.6076 | -36.1008
26.3277 | -35.608 | -36.1061
26.3338 | -35.6083 | -36.1114
26.3399 | -35.6086 | -36.1167
26.346 | -35.609 | -36.1218
26.3521 | -35.6093 | -36.1273
26.3582 | -35.6097 | -36.1327
26.3643 | -35.61 | -36.1379
26.3704 | -35.6104 | -36.1432
26.3765 | -35.6107 | -36.1485
26.3826 | -35.611 | -36.1537
26.3887 | -35.6114 | -36.1591
26.3948 | -35.6117 | -36.1644
26.4009 | -35.6121 | -36.1697
26.407 | -35.6124 | -36.175
26.4131 | -35.6128 | -36.1804
26.4192 | -35.6131 | -36.1855
26.4253 | -35.6135 | -36.1908
26.4314 | -35.6138 | -36.196
26.4376 | -35.6141 | -36.2013
26.4437 | -35.6145 | -36.2066
26.4498 | -35.6148 | -36.212
26.4559 | -35.6151 | -36.2172
26.462 | -35.6154 | -36.2226
26.4681 | -35.6158 | -36.2279
26.4742 | -35.6161 | -36.2332
26.4803 | -35.6165 | -36.2386
26.4864 | -35.6168 | -36.2438
26.4925 | -35.6171 | -36.249
26.4986 | -35.6175 | -36.2543
26.5047 | -35.6178 | -36.2596
26.5108 | -35.6182 | -36.2649
26.5169 | -35.6185 | -36.2702
26.523 | -35.6188 | -36.2752
26.5291 | -35.6192 | -36.2805
26.5352 | -35.6195 | -36.2857
26.5413 | -35.6199 | -36.2911
26.5474 | -35.6202 | -36.2964
26.5535 | -35.6206 | -36.3017
26.5596 | -35.6209 | -36.307
26.5657 | -35.6212 | -36.3122
26.5718 | -35.6216 | -36.3175
26.578 | -35.6219 | -36.3228
26.5841 | -35.6222 | -36.328
26.5902 | -35.6226 | -36.3333
26.5963 | -35.6229 | -36.3384
26.6024 | -35.6232 | -36.3437
26.6085 | -35.6236 | -36.3491
26.6146 | -35.6239 | -36.3545
26.6207 | -35.6242 | -36.3599
26.6268 | -35.6246 | -36.3652
26.6329 | -35.6249 | -36.3704
26.639 | -35.6252 | -36.3757
26.6451 | -35.6255 | -36.381
26.6512 | -35.6259 | -36.3862
26.6573 | -35.6262 | -36.3915
26.6634 | -35.6265 | -36.3966
26.6695 | -35.6269 | -36.4021
26.6756 | -35.6272 | -36.4074
26.6817 | -35.6275 | -36.4126
26.6878 | -35.6279 | -36.418
26.6939 | -35.6282 | -36.4232
26.7 | -35.6285 | -36.4282
26.7061 | -35.6289 | -36.4335
26.7122 | -35.6292 | -36.4388
26.7183 | -35.6295 | -36.444
26.7245 | -35.6299 | -36.4494
26.7306 | -35.6302 | -36.4545
26.7367 | -35.6305 | -36.4596
26.7428 | -35.6309 | -36.4649
26.7489 | -35.6312 | -36.4701
26.755 | -35.6316 | -36.4753
26.7611 | -35.6319 | -36.4806
26.7672 | -35.6322 | -36.4859
26.7733 | -35.6325 | -36.491
26.7794 | -35.6329 | -36.4963
26.7855 | -35.6332 | -36.5016
26.7916 | -35.6336 | -36.5068
26.7977 | -35.6339 | -36.5121
26.8038 | -35.6342 | -36.5173
26.8099 | -35.6345 | -36.5224
26.816 | -35.6349 | -36.5277
26.8221 | -35.6352 | -36.5329
26.8282 | -35.6355 | -36.538
26.8343 | -35.6358 | -36.5433
26.8404 | -35.6362 | -36.5484
26.8465 | -35.6365 | -36.5537
26.8526 | -35.6368 | -36.5588
26.8587 | -35.6372 | -36.5641
26.8649 | -35.6375 | -36.5693
26.871 | -35.6378 | -36.5746
26.8771 | -35.6382 | -36.5797
26.8832 | -35.6385 | -36.585
26.8893 | -35.6388 | -36.5902
26.8954 | -35.6391 | -36.5954
26.9015 | -35.6394 | -36.6006
26.9076 | -35.6398 | -36.6059
26.9137 | -35.6401 | -36.6111
26.9198 | -35.6404 | -36.6164
26.9259 | -35.6408 | -36.6217
26.932 | -35.6411 | -36.627
26.9381 | -35.6414 | -36.6323
26.9442 | -35.6417 | -36.6374
26.9503 | -35.6421 | -36.6426
26.9564 | -35.6424 | -36.6477
26.9625 | -35.6427 | -36.6529
26.9686 | -35.643 | -36.6581
26.9747 | -35.6434 | -36.6633
26.9808 | -35.6437 | -36.6684
26.9869 | -35.644 | -36.6736
26.993 | -35.6444 | -36.6789
26.9991 | -35.6447 | -36.6841
27.0052 | -35.645 | -36.6893
27.0114 | -35.6453 | -36.6947
27.0175 | -35.6456 | -36.6997
27.0236 | -35.646 | -36.7049
27.0297 | -35.6463 | -36.7101
27.0358 | -35.6466 | -36.7153
27.0419 | -35.6469 | -36.7205
27.048 | -35.6473 | -36.7256
27.0541 | -35.6476 | -36.7307
27.0602 | -35.6479 | -36.7358
27.0663 | -35.6483 | -36.741
27.0724 | -35.6486 | -36.7462
27.0785 | -35.6489 | -36.7515
27.0846 | -35.6492 | -36.7564
27.0907 | -35.6495 | -36.7616
27.0968 | -35.6498 | -36.7669
27.1029 | -35.6502 | -36.7721
27.109 | -35.6505 | -36.7772
27.1151 | -35.6508 | -36.7825
27.1212 | -35.6511 | -36.7876
27.1273 | -35.6515 | -36.7926
27.1334 | -35.6518 | -36.7978
27.1395 | -35.6521 | -36.803
27.1456 | -35.6525 | -36.8083
27.1518 | -35.6528 | -36.8136
27.1579 | -35.6531 | -36.8186
27.164 | -35.6534 | -36.8238
27.1701 | -35.6537 | -36.8289
27.1762 | -35.654 | -36.8342
27.1823 | -35.6543 | -36.8393
27.1884 | -35.6547 | -36.8444
27.1945 | -35.655 | -36.8495
27.2006 | -35.6553 | -36.8547
27.2067 | -35.6556 | -36.8598
27.2128 | -35.6559 | -36.8651
27.2189 | -35.6563 | -36.8702
27.225 | -35.6566 | -36.8755
27.2311 | -35.6569 | -36.8806
27.2372 | -35.6572 | -36.8857
27.2433 | -35.6575 | -36.8909
27.2494 | -35.6579 | -36.8961
27.2555 | -35.6582 | -36.9013
27.2616 | -35.6585 | -36.9064
27.2677 | -35.6588 | -36.9116
27.2738 | -35.6591 | -36.9167
27.2799 | -35.6595 | -36.922
27.286 | -35.6598 | -36.9271
27.2921 | -35.6601 | -36.9323
27.2983 | -35.6604 | -36.9373
27.3044 | -35.6607 | -36.9426
27.3105 | -35.661 | -36.9477
27.3166 | -35.6613 | -36.9528
27.3227 | -35.6617 | -36.9581
27.3288 | -35.662 | -36.9633
27.3349 | -35.6623 | -36.9683
27.341 | -35.6626 | -36.9735
27.3471 | -35.6629 | -36.9785
27.3532 | -35.6632 | -36.9837
27.3593 | -35.6635 | -36.9889
27.3654 | -35.6638 | -36.9941
27.3715 | -35.6641 | -36.999
27.3776 | -35.6645 | -37.0043
27.3837 | -35.6648 | -37.0094
27.3898 | -35.6651 | -37.0145
27.3959 | -35.6654 | -37.0197
27.402 | -35.6657 | -37.0246
27.4081 | -35.666 | -37.0297
27.4142 | -35.6664 | -37.0349
27.4203 | -35.6667 | -37.04
27.4264 | -35.667 | -37.0452
27.4325 | -35.6673 | -37.0503
27.4387 | -35.6676 | -37.0554
27.4448 | -35.6679 | -37.0604
27.4509 | -35.6682 | -37.0656
27.457 | -35.6686 | -37.0708
27.4631 | -35.6689 | -37.076
27.4692 | -35.6692 | -37.0811
27.4753 | -35.6695 | -37.0861
27.4814 | -35.6698 | -37.0912
27.4875 | -35.6701 | -37.0963
27.4936 | -35.6704 | -37.1014
27.4997 | -35.6707 | -37.1065
27.5058 | -35.6711 | -37.1117
27.5119 | -35.6714 | -37.1168
27.518 | -35.6717 | -37.1219
27.5241 | -35.672 | -37.127
27.5302 | -35.6723 | -37.1322
27.5363 | -35.6726 | -37.1374
27.5424 | -35.673 | -37.1423
27.5485 | -35.6733 | -37.1475
27.5546 | -35.6736 | -37.1526
27.5607 | -35.6739 | -37.1577
27.5668 | -35.6742 | -37.1628
27.5729 | -35.6745 | -37.1678
27.5791 | -35.6748 | -37.1728
27.5852 | -35.6751 | -37.1779
27.5913 | -35.6755 | -37.1829
27.5974 | -35.6758 | -37.188
27.6035 | -35.6761 | -37.1931
27.6096 | -35.6764 | -37.1983
27.6157 | -35.6766 | -37.2032
27.6218 | -35.677 | -37.2082
27.6279 | -35.6773 | -37.2133
27.634 | -35.6776 | -37.2183
27.6401 | -35.6779 | -37.2234
27.6462 | -35.6782 | -37.2286
27.6523 | -35.6785 | -37.2335
27.6584 | -35.6788 | -37.2386
27.6645 | -35.6792 | -37.2438
27.6706 | -35.6795 | -37.249
27.6767 | -35.6798 | -37.2541
27.6828 | -35.68 | -37.2589
27.6889 | -35.6803 | -37.264
27.695 | -35.6806 | -37.269
27.7011 | -35.6809 | -37.2741
27.7072 | -35.6812 | -37.2792
27.7133 | -35.6815 | -37.2844
27.7194 | -35.6819 | -37.2895
27.7256 | -35.6822 | -37.2946
27.7317 | -35.6825 | -37.2997
27.7378 | -35.6828 | -37.3047
27.7439 | -35.6831 | -37.3097
27.75 | -35.6834 | -37.3149
27.7561 | -35.6837 | -37.3198
27.7622 | -35.684 | -37.3248
27.7683 | -35.6843 | -37.3299
27.7744 | -35.6847 | -37.335
27.7805 | -35.685 | -37.34
27.7866 | -35.6853 | -37.3451
27.7927 | -35.6856 | -37.3499
27.7988 | -35.6859 | -37.3549
27.8049 | -35.6862 | -37.36
27.811 | -35.6865 | -37.365
27.8171 | -35.6868 | -37.3701
27.8232 | -35.6871 | -37.3752
27.8293 | -35.6874 | -37.3803
27.8354 | -35.6877 | -37.3853
27.8415 | -35.688 | -37.3903
27.8476 | -35.6883 | -37.3954
27.8537 | -35.6886 | -37.4006
27.8598 | -35.6889 | -37.4054
27.866 | -35.6892 | -37.4105
27.8721 | -35.6895 | -37.4155
27.8782 | -35.6898 | -37.4207
27.8843 | -35.6901 | -37.4258
27.8904 | -35.6904 | -37.4309
27.8965 | -35.6907 | -37.4357
27.9026 | -35.691 | -37.4408
27.9087 | -35.6913 | -37.4459
27.9148 | -35.6916 | -37.4509
27.9209 | -35.6919 | -37.456
27.927 | -35.6922 | -37.461
27.9331 | -35.6925 | -37.4662
27.9392 | -35.6928 | -37.4712
27.9453 | -35.6931 | -37.4763
27.9514 | -35.6934 | -37.4814
27.9575 | -35.6937 | -37.4865
27.9636 | -35.694 | -37.4916
27.9697 | -35.6943 | -37.4965
27.9758 | -35.6946 | -37.5017
27.9819 | -35.6949 | -37.5068
27.988 | -35.6952 | -37.5119
27.9941 | -35.6955 | -37.5169
28.0002 | -35.6958 | -37.5216
28.0063 | -35.6961 | -37.5265
28.0125 | -35.6964 | -37.5315
28.0186 | -35.6967 | -37.5367
28.0247 | -35.697 | -37.5418
28.0308 | -35.6973 | -37.5468
28.0369 | -35.6976 | -37.5517
28.043 | -35.6978 | -37.5567
28.0491 | -35.6981 | -37.5618
28.0552 | -35.6984 | -37.5669
28.0613 | -35.6987 | -37.5719
28.0674 | -35.699 | -37.5769
28.0735 | -35.6993 | -37.5817

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4567 | -42.3941
21.4908 | -35.5076 | -34.6524
16.4948 | -34.8992 | -27.3587
12.6602 | -34.5418 | -21.151
9.7171 | -34.3311 | -16.2895
7.4582 | -34.1845 | -12.3475
5.7244 | -34.0867 | -9.153
4.3936 | -34.0179 | -6.5762
3.3722 | -34.0076 | -4.5451
2.5883 | -34.0046 | -3.0068
1.9866 | -34.0049 | -1.779
1.5248 | -33.9994 | -0.79669
1.1703 | -34.0066 | -0.15352
1 | -34.0177 | 0.16172
0.89822 | -34.0072 | 0.43179
0.68941 | -33.9881 | 1.0073
0.52919 | -33.9973 | 1.4511
0.40612 | -34.002 | 2.0503
0.31171 | -33.9917 | 2.5024
0.23928 | -33.9986 | 3.2209
0.18358 | -34.0013 | 3.6391
0.14091 | -34.0246 | 4.5152
0.10819 | -34.0442 | 5.6707
0.083017 | -34.0659 | 7.1434
0.063705 | -34.1048 | 9.0133
0.048923 | -34.1797 | 11.5064
0.037479 | -34.2902 | 14.6929
0.028801 | -34.4693 | 18.709
0.022078 | -34.7784 | 24.024
0.017023 | -35.262 | 30.4049
0.013018 | -36.0627 | 38.057
0.010014 | -37.1703 | 45.9495

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003673 | -12.7339
0.007347 | -29.6025
0.01102 | -37.4095
0.014693 | -41.798
0.018366 | -43.6601
0.02204 | -45.5003
0.025713 | -47.5159
0.029386 | -49.1538
0.033059 | -50.4515
0.036733 | -51.1426
0.040406 | -52.7236
0.044079 | -54.4171
0.047752 | -55.1939
0.051426 | -55.0755
0.055099 | -55.3214
0.058772 | -57.3818
0.062445 | -57.9634
0.066119 | -58.7758
0.069792 | -59.2686
0.073465 | -60.1698
0.077139 | -61.7368
0.080812 | -61.8577
0.084485 | -61.1426
0.088158 | -62.1078
0.091832 | -62.3014
0.095505 | -61.9937
0.099178 | -62.9973
0.10285 | -64.4928
0.10652 | -64.7161
0.1102 | -64.4006
0.11387 | -63.8965
0.11754 | -64.8625
0.12122 | -65.1904
0.12489 | -66.5099
0.12856 | -67.7327
0.13224 | -67.3682
0.13591 | -66.6126
0.13958 | -67.1857
0.14326 | -66.9989
0.14693 | -67.0322
0.1506 | -67.5709
0.15428 | -68.3841
0.15795 | -69.0351
0.16162 | -68.813
0.1653 | -68.9752
0.16897 | -69.4821
0.17264 | -69.4054
0.17632 | -69.9191
0.17999 | -70.619
0.18366 | -71.1415
0.18734 | -70.3976
0.19101 | -71.2267
0.19468 | -71.3054
0.19836 | -71.8016
0.20203 | -72.0039
0.2057 | -72.13
0.20938 | -71.9331
0.21305 | -71.4166
0.21672 | -71.44
0.2204 | -71.62
0.22407 | -72.7066
0.22774 | -73.1728
0.23142 | -73.3419
0.23509 | -74.3425
0.23876 | -74.3538
0.24244 | -74.4955
0.24611 | -73.9826
0.24978 | -74.1845
0.25345 | -74.8454
0.25713 | -74.3012
0.2608 | -73.9077
0.26448 | -74.2506
0.26815 | -74.3252
0.27182 | -75.0115
0.27549 | -74.5495
0.27917 | -74.7892
0.28284 | -75.7626
0.28652 | -75.1894
0.29019 | -74.7595
0.29386 | -75.3304
0.29753 | -74.937
0.30121 | -75.7422
0.30488 | -75.5751
0.30855 | -75.8564
0.31223 | -76.2419
0.3159 | -75.7922
0.31957 | -75.08
0.32325 | -75.9475
0.32692 | -77.2082
0.33059 | -77.6261
0.33427 | -77.8506
0.33794 | -78.7457
0.34161 | -77.5101
0.34529 | -76.7352
0.34896 | -76.8053
0.35263 | -76.6333
0.35631 | -76.9916
0.35998 | -77.4236
0.36365 | -77.4527
0.36733 | -77.6345
0.371 | -77.658
0.37467 | -78.5985
0.37835 | -78.1512
0.38202 | -79.0564
0.38569 | -78.67
0.38937 | -78.7622
0.39304 | -78.2624
0.39671 | -78.191
0.40039 | -77.4537
0.40406 | -77.6848
0.40773 | -78.2403
0.41141 | -78.1808
0.41508 | -78.4132
0.41875 | -79.5371
0.42242 | -79.4652
0.4261 | -79.8837
0.42977 | -78.9598
0.43345 | -79.3768
0.43712 | -79.8503
0.44079 | -78.8848
0.44446 | -79.3563
0.44814 | -80.0841
0.45181 | -80.1737
0.45548 | -80.1184
0.45916 | -79.8917
0.46283 | -81.1638
0.4665 | -81.5027
0.47018 | -80.3179
0.47385 | -79.5736
0.47752 | -80.7362
0.4812 | -80.4785
0.48487 | -80.1875
0.48854 | -79.5844
0.49222 | -80.1467
0.49589 | -80.1681
0.49956 | -80.3722
0.50324 | -80.5123
0.50691 | -80.2407
0.51058 | -80.7473
0.51426 | -81.3364
0.51793 | -81.1435
0.5216 | -80.8832
0.52528 | -81.0505
0.52895 | -81.2261
0.53262 | -81.1761
0.5363 | -81.4219
0.53997 | -81.8996
0.54364 | -82.0333
0.54732 | -81.9542
0.55099 | -82.0922
0.55466 | -82.3435
0.55834 | -82.317
0.56201 | -82.2652
0.56568 | -82.7215
0.56936 | -82.8628
0.57303 | -82.5637
0.5767 | -82.3465
0.58038 | -82.7938
0.58405 | -83.4088
0.58772 | -83.075
0.5914 | -82.4372
0.59507 | -82.2983
0.59874 | -82.6313
0.60242 | -82.9214
0.60609 | -83.0111
0.60976 | -83.2912
0.61343 | -83.6288
0.61711 | -83.5556
0.62078 | -83.1344
0.62445 | -82.9676
0.62813 | -83.2086
0.6318 | -83.3566
0.63548 | -83.2059
0.63915 | -83.255
0.64282 | -83.4398
0.64649 | -83.5588
0.65017 | -83.156
0.65384 | -82.7645
0.65751 | -82.9701
0.66119 | -83.5193
0.66486 | -83.5042
0.66853 | -82.8745
0.67221 | -82.84
0.67588 | -83.3158
0.67955 | -83.8722
0.68323 | -84.2834
0.6869 | -84.3134
0.69057 | -83.958
0.69425 | -83.7647
0.69792 | -83.8275
0.70159 | -83.631
0.70527 | -83.7159
0.70894 | -84.3541
0.71261 | -85.087
0.71629 | -85.3568
0.71996 | -85.2921
0.72363 | -84.8266
0.72731 | -84.2251
0.73098 | -84.2401
0.73465 | -84.4565
0.73833 | -84.4403
0.742 | -84.7242
0.74567 | -85.149
0.74935 | -85.0883
0.75302 | -84.9902
0.75669 | -85.1597
0.76036 | -84.9302
0.76404 | -84.2796
0.76771 | -83.9775
0.77138 | -84.3491
0.77506 | -84.9435
0.77873 | -85.0692
0.78241 | -85.1651
0.78608 | -85.2838
0.78975 | -85.4446
0.79343 | -85.653
0.7971 | -86.0347
0.80077 | -86.1984
0.80444 | -85.8502
0.80812 | -85.7688
0.81179 | -86.1672
0.81546 | -86.1039
0.81914 | -85.4272
0.82281 | -84.8256
0.82648 | -85.0889
0.83016 | -85.7079
0.83383 | -85.7277
0.8375 | -85.7208
0.84118 | -85.976
0.84485 | -85.7687
0.84852 | -85.4089
0.8522 | -85.3546
0.85587 | -85.5356
0.85954 | -85.6478
0.86322 | -85.65
0.86689 | -85.9937
0.87056 | -86.0166
0.87424 | -85.6011
0.87791 | -85.5008
0.88158 | -85.7562
0.88526 | -85.9374
0.88893 | -85.9777
0.8926 | -86.0319
0.89628 | -86.4584
0.89995 | -86.8305
0.90362 | -86.8903
0.9073 | -86.921
0.91097 | -86.4735
0.91464 | -86.1439
0.91832 | -86.3323
0.92199 | -86.4539
0.92566 | -86.5236
0.92934 | -86.611
0.93301 | -86.9482
0.93668 | -87.144
0.94036 | -86.8412
0.94403 | -86.6382
0.9477 | -86.7604
0.95137 | -87.1757
0.95505 | -87.663
0.95872 | -88.0204
0.9624 | -88.0678
0.96607 | -87.6405
0.96974 | -87.4545
0.97342 | -87.6408
0.97709 | -87.9369
0.98076 | -88.0417
0.98443 | -88.1613
0.98811 | -87.9529
0.99178 | -87.651
0.99545 | -87.7639
0.99913 | -87.6647
1.0028 | -87.1751
1.0065 | -87.1593
1.0101 | -87.4914
1.0138 | -87.5363
1.0175 | -87.3512
1.0212 | -87.3823
1.0248 | -87.5492
1.0285 | -87.7861
1.0322 | -88.1006
1.0359 | -87.9691
1.0395 | -87.5227
1.0432 | -87.1474
1.0469 | -86.886
1.0506 | -87.1537
1.0542 | -87.7314
1.0579 | -88.2738
1.0616 | -88.6015
1.0652 | -88.5075
1.0689 | -88.1548
1.0726 | -87.9678
1.0763 | -88.1305
1.0799 | -88.3746
1.0836 | -88.4609
1.0873 | -88.4691
1.091 | -88.4998
1.0946 | -88.5501
1.0983 | -88.5975
1.102 | -88.704
1.1057 | -88.9587
1.1093 | -89.3092
1.113 | -89.56
1.1167 | -89.4559
1.1203 | -88.9791
1.124 | -88.493
1.1277 | -88.2933
1.1314 | -88.2852
1.135 | -88.2124
1.1387 | -88.08
1.1424 | -88.0193
1.1461 | -88.0665
1.1497 | -88.1134
1.1534 | -88.0869
1.1571 | -88.1275
1.1608 | -88.3899
1.1644 | -88.8
1.1681 | -89.1003
1.1718 | -89.1671
1.1754 | -89.2173
1.1791 | -89.3345
1.1828 | -89.495
1.1865 | -89.7542
1.1901 | -90.063
1.1938 | -90.2924
1.1975 | -90.2972
1.2012 | -90.0537
1.2048 | -89.7754
1.2085 | -89.6016
1.2122 | -89.5236
1.2159 | -89.3861
1.2195 | -89.0105
1.2232 | -88.4949
1.2269 | -88.1492
1.2305 | -88.2183
1.2342 | -88.5766
1.2379 | -88.8085
1.2416 | -88.7958
1.2452 | -88.7393
1.2489 | -88.7595
1.2526 | -88.8767
1.2563 | -89.1123
1.2599 | -89.4981
1.2636 | -89.998
1.2673 | -90.5293
1.2709 | -90.9547
1.2746 | -91.0722
1.2783 | -90.8542
1.282 | -90.5099
1.2856 | -90.206
1.2893 | -89.8393
1.293 | -89.2836
1.2967 | -88.7196
1.3003 | -88.4386
1.304 | -88.5049
1.3077 | -88.7269
1.3114 | -88.9292
1.315 | -89.1461
1.3187 | -89.4428
1.3224 | -89.7813
1.326 | -90.0764
1.3297 | -90.2232
1.3334 | -90.2292
1.3371 | -90.1513
1.3407 | -89.8844
1.3444 | -89.456
1.3481 | -89.1236
1.3518 | -89.0743
1.3554 | -89.2999
1.3591 | -89.5755
1.3628 | -89.6616
1.3665 | -89.5622
1.3701 | -89.5318
1.3738 | -89.8204
1.3775 | -90.3056
1.3811 | -90.5231
1.3848 | -90.3442
1.3885 | -90.0291
1.3922 | -89.7143
1.3958 | -89.433
1.3995 | -89.3221
1.4032 | -89.5897
1.4069 | -90.2088
1.4105 | -90.7412
1.4142 | -90.8073
1.4179 | -90.5215
1.4216 | -90.1371
1.4252 | -89.8457
1.4289 | -89.8036
1.4326 | -90.0136
1.4362 | -90.2927
1.4399 | -90.4894
1.4436 | -90.5174
1.4473 | -90.3952
1.4509 | -90.4088
1.4546 | -90.7263
1.4583 | -91.0995
1.462 | -91.2102
1.4656 | -91.058
1.4693 | -90.9069
1.473 | -90.8681
1.4767 | -90.7843
1.4803 | -90.6929
1.484 | -90.8103
1.4877 | -91.0237
1.4913 | -91.0518
1.495 | -90.9102
1.4987 | -90.7291
1.5024 | -90.6545
1.506 | -90.8661
1.5097 | -91.2813
1.5134 | -91.5657
1.5171 | -91.594
1.5207 | -91.534
1.5244 | -91.441
1.5281 | -91.3359
1.5318 | -91.3725
1.5354 | -91.5851
1.5391 | -91.8613
1.5428 | -92.1184
1.5464 | -92.2281
1.5501 | -92.0954
1.5538 | -91.8203
1.5575 | -91.5807
1.5611 | -91.4755
1.5648 | -91.429
1.5685 | -91.2962
1.5722 | -91.2035
1.5758 | -91.3526
1.5795 | -91.5711
1.5832 | -91.5764
1.5869 | -91.5524
1.5905 | -91.6134
1.5942 | -91.7099
1.5979 | -91.762
1.6015 | -91.7049
1.6052 | -91.5241
1.6089 | -91.2599
1.6126 | -90.9832
1.6162 | -90.7769
1.6199 | -90.7267
1.6236 | -90.8505
1.6273 | -91.0569
1.6309 | -91.2281
1.6346 | -91.3608
1.6383 | -91.5526
1.6419 | -91.852
1.6456 | -92.1967
1.6493 | -92.4931
1.653 | -92.6856
1.6566 | -92.743
1.6603 | -92.6733
1.664 | -92.5785
1.6677 | -92.5606
1.6713 | -92.6413
1.675 | -92.73
1.6787 | -92.7165
1.6824 | -92.5803
1.686 | -92.3975
1.6897 | -92.2672
1.6934 | -92.187
1.697 | -92.1114
1.7007 | -92.0517
1.7044 | -92.0482
1.7081 | -92.0896
1.7117 | -92.1089
1.7154 | -92.0605
1.7191 | -91.9629
1.7228 | -91.8824
1.7264 | -91.8831
1.7301 | -91.9834
1.7338 | -92.1196
1.7375 | -92.162
1.7411 | -92.0635
1.7448 | -91.9414
1.7485 | -91.963
1.7521 | -92.1459
1.7558 | -92.3382
1.7595 | -92.4028
1.7632 | -92.3723
1.7668 | -92.3058
1.7705 | -92.2049
1.7742 | -92.1251
1.7779 | -92.118
1.7815 | -92.1334
1.7852 | -92.1126
1.7889 | -92.0784
1.7926 | -92.0857
1.7962 | -92.1267
1.7999 | -92.1959
1.8036 | -92.3007
1.8072 | -92.442
1.8109 | -92.6286
1.8146 | -92.8205
1.8183 | -92.9324
1.8219 | -92.9309
1.8256 | -92.8327
1.8293 | -92.6894
1.833 | -92.5666
1.8366 | -92.5311
1.8403 | -92.6119
1.844 | -92.7422
1.8477 | -92.793
1.8513 | -92.7086
1.855 | -92.5539
1.8587 | -92.4233
1.8623 | -92.3697
1.866 | -92.4133
1.8697 | -92.5403
1.8734 | -92.7041
1.877 | -92.8357
1.8807 | -92.891
1.8844 | -92.8722
1.8881 | -92.8235
1.8917 | -92.7568
1.8954 | -92.6598
1.8991 | -92.5987
1.9027 | -92.6748
1.9064 | -92.8587
1.9101 | -93.046
1.9138 | -93.1655
1.9174 | -93.2366
1.9211 | -93.2977
1.9248 | -93.3273
1.9285 | -93.3072
1.9321 | -93.2477
1.9358 | -93.1801
1.9395 | -93.1024
1.9432 | -92.9553
1.9468 | -92.7707
1.9505 | -92.6494
1.9542 | -92.6429
1.9578 | -92.7164
1.9615 | -92.7999
1.9652 | -92.8892
1.9689 | -93.0091
1.9725 | -93.1592
1.9762 | -93.2524
1.9799 | -93.2444
1.9836 | -93.2167
1.9872 | -93.2873
1.9909 | -93.4222
1.9946 | -93.488
1.9983 | -93.4319
2.0019 | -93.2961
2.0056 | -93.1514
2.0093 | -93.0751
2.0129 | -93.0809
2.0166 | -93.1111
2.0203 | -93.0946
2.024 | -93.0803
2.0276 | -93.1519
2.0313 | -93.3453
2.035 | -93.5958
2.0387 | -93.7468
2.0423 | -93.7248
2.046 | -93.627
2.0497 | -93.5704
2.0534 | -93.5481
2.057 | -93.4975
2.0607 | -93.4464
2.0644 | -93.4678
2.068 | -93.5822
2.0717 | -93.7174
2.0754 | -93.7329
2.0791 | -93.5726
2.0827 | -93.3095
2.0864 | -93.0946
2.0901 | -93.0582
2.0938 | -93.2338
2.0974 | -93.5483
2.1011 | -93.8365
2.1048 | -93.9718
2.1085 | -93.9855
2.1121 | -93.9784
2.1158 | -93.9998
2.1195 | -93.8888
2.1231 | -93.7396
2.1268 | -93.5362
2.1305 | -93.3555
2.1342 | -93.2717
2.1378 | -93.3139
2.1415 | -93.4571
2.1452 | -93.6509
2.1489 | -93.845
2.1525 | -93.9992
2.1562 | -94.0889
2.1599 | -94.1149
2.1636 | -94.1148
2.1672 | -94.1142
2.1709 | -94.1193
2.1746 | -94.1351
2.1782 | -94.1781
2.1819 | -94.2137
2.1856 | -94.1794
2.1893 | -94.0594
2.1929 | -93.8956
2.1966 | -93.7706
2.2003 | -93.7373
2.204 | -93.8066
2.2076 | -93.9429
2.2113 | -94.0598
2.215 | -94.0962
2.2187 | -94.0438
2.2223 | -93.9603
2.226 | -93.9207
2.2297 | -93.9786
2.2333 | -94.1506
2.237 | -94.3532
2.2407 | -94.4697
2.2444 | -94.3981
2.248 | -94.1536
2.2517 | -93.8317
2.2554 | -93.5265
2.2591 | -93.3226
2.2627 | -93.284
2.2664 | -93.4305
2.2701 | -93.7273
2.2738 | -94.0967
2.2774 | -94.4355
2.2811 | -94.6408
2.2848 | -94.662
2.2884 | -94.537
2.2921 | -94.3462
2.2958 | -94.1619
2.2995 | -94.0258
2.3031 | -93.9498
2.3068 | -93.9489
2.3105 | -94.0244
2.3142 | -94.1608
2.3178 | -94.306
2.3215 | -94.4036
2.3252 | -94.4554
2.3288 | -94.4742
2.3325 | -94.473
2.3362 | -94.471
2.3399 | -94.4834
2.3435 | -94.5377
2.3472 | -94.6147
2.3509 | -94.6804
2.3546 | -94.686
2.3582 | -94.6488
2.3619 | -94.5968
2.3656 | -94.5525
2.3693 | -94.5426
2.3729 | -94.5739
2.3766 | -94.6126
2.3803 | -94.6577
2.3839 | -94.6744
2.3876 | -94.6683
2.3913 | -94.6144
2.395 | -94.529
2.3986 | -94.4677
2.4023 | -94.4985
2.406 | -94.6334
2.4097 | -94.8182
2.4133 | -94.9763
2.417 | -95.0804
2.4207 | -95.1299
2.4244 | -95.1749
2.428 | -95.2023
2.4317 | -95.2107
2.4354 | -95.1698
2.439 | -95.0812
2.4427 | -94.9597
2.4464 | -94.8096
2.4501 | -94.6363
2.4537 | -94.4648
2.4574 | -94.3532
2.4611 | -94.3547
2.4648 | -94.4582
2.4684 | -94.6187
2.4721 | -94.7645
2.4758 | -94.8499
2.4795 | -94.8467
2.4831 | -94.7862
2.4868 | -94.7042
2.4905 | -94.6368
2.4941 | -94.6166
2.4978 | -94.6507
2.5015 | -94.7326
2.5052 | -94.851
2.5088 | -94.9529
2.5125 | -95.0123
2.5162 | -95.0343
2.5199 | -95.045
2.5235 | -95.0546
2.5272 | -95.048
2.5309 | -94.9981
2.5346 | -94.9213
2.5382 | -94.8526
2.5419 | -94.8165
2.5456 | -94.8043
2.5492 | -94.8233
2.5529 | -94.8522
2.5566 | -94.8653
2.5603 | -94.8538
2.5639 | -94.8263
2.5676 | -94.8173
2.5713 | -94.8565
2.575 | -94.9446
2.5786 | -95.0691
2.5823 | -95.2043
2.586 | -95.3037
2.5897 | -95.3302
2.5933 | -95.2858
2.597 | -95.2108
2.6007 | -95.1542
2.6043 | -95.1433
2.608 | -95.1684
2.6117 | -95.2067
2.6154 | -95.2652
2.619 | -95.3658
2.6227 | -95.5193
2.6264 | -95.6954
2.6301 | -95.8549
2.6337 | -95.9477
2.6374 | -95.9408
2.6411 | -95.8579
2.6448 | -95.7553
2.6484 | -95.6864
2.6521 | -95.6303
2.6558 | -95.6098
2.6594 | -95.625
2.6631 | -95.6546
2.6668 | -95.675
2.6705 | -95.6658
2.6741 | -95.6246
2.6778 | -95.5804
2.6815 | -95.5556
2.6852 | -95.5537
2.6888 | -95.5889
2.6925 | -95.6646
2.6962 | -95.7717
2.6998 | -95.8633
2.7035 | -95.8696
2.7072 | -95.7884
2.7109 | -95.6571
2.7145 | -95.5501
2.7182 | -95.5189
2.7219 | -95.5631
2.7256 | -95.649
2.7292 | -95.7478
2.7329 | -95.8458
2.7366 | -95.9243
2.7403 | -95.9832
2.7439 | -96.0115
2.7476 | -96.0125
2.7513 | -95.9945
2.7549 | -95.9494
2.7586 | -95.8902
2.7623 | -95.8342
2.766 | -95.7988
2.7696 | -95.779
2.7733 | -95.7719
2.777 | -95.7621
2.7807 | -95.741
2.7843 | -95.7004
2.788 | -95.6265
2.7917 | -95.519
2.7954 | -95.4106
2.799 | -95.3411
2.8027 | -95.3313
2.8064 | -95.3801
2.81 | -95.4463
2.8137 | -95.5023
2.8174 | -95.5423
2.8211 | -95.5778
2.8247 | -95.621
2.8284 | -95.6821
2.8321 | -95.7723
2.8358 | -95.8794
2.8394 | -95.9977
2.8431 | -96.0987
2.8468 | -96.1758
2.8505 | -96.2251
2.8541 | -96.2296
2.8578 | -96.1747
2.8615 | -96.0639
2.8651 | -95.9104
2.8688 | -95.75
2.8725 | -95.6286
2.8762 | -95.5705
2.8798 | -95.5665
2.8835 | -95.6252
2.8872 | -95.7264
2.8909 | -95.8618
2.8945 | -96.0096
2.8982 | -96.1655
2.9019 | -96.338
2.9056 | -96.5119
2.9092 | -96.6609
2.9129 | -96.763
2.9166 | -96.812
2.9202 | -96.812
2.9239 | -96.7911
2.9276 | -96.7781
2.9313 | -96.7768
2.9349 | -96.7887
2.9386 | -96.7739
2.9423 | -96.7024
2.946 | -96.5618
2.9496 | -96.3764
2.9533 | -96.1988
2.957 | -96.0783
2.9606 | -96.0397
2.9643 | -96.08
2.968 | -96.185
2.9717 | -96.3042
2.9753 | -96.3952
2.979 | -96.431
2.9827 | -96.4022
2.9864 | -96.33
2.99 | -96.2419
2.9937 | -96.1772
2.9974 | -96.1413
3.0011 | -96.1407
3.0047 | -96.1673
3.0084 | -96.2146
3.0121 | -96.2651
3.0157 | -96.2903
3.0194 | -96.28
3.0231 | -96.2428
3.0268 | -96.1907
3.0304 | -96.1416
3.0341 | -96.0988
3.0378 | -96.0644
3.0415 | -96.0305
3.0451 | -96.0092
3.0488 | -96.0023
3.0525 | -96.0299
3.0562 | -96.0833
3.0598 | -96.1396
3.0635 | -96.1899
3.0672 | -96.2271
3.0708 | -96.2548
3.0745 | -96.2971
3.0782 | -96.3618
3.0819 | -96.4503
3.0855 | -96.5609
3.0892 | -96.6841
3.0929 | -96.7907
3.0966 | -96.8717
3.1002 | -96.9214
3.1039 | -96.9586
3.1076 | -96.9869
3.1113 | -96.9925
3.1149 | -96.9778
3.1186 | -96.9433
3.1223 | -96.9131
3.1259 | -96.9092
3.1296 | -96.9243
3.1333 | -96.9553
3.137 | -96.9883
3.1406 | -97.0111
3.1443 | -97.0086
3.148 | -97.0081
3.1517 | -97.0452
3.1553 | -97.12
3.159 | -97.2059
3.1627 | -97.2664
3.1664 | -97.2852
3.17 | -97.2437
3.1737 | -97.1698
3.1774 | -97.0781
3.181 | -97.0066
3.1847 | -96.9365
3.1884 | -96.8719
3.1921 | -96.819
3.1957 | -96.7949
3.1994 | -96.8037
3.2031 | -96.8547
3.2068 | -96.9548
3.2104 | -97.1085
3.2141 | -97.2956
3.2178 | -97.4977
3.2215 | -97.7053
3.2251 | -97.8688
3.2288 | -97.9636
3.2325 | -97.9706
3.2361 | -97.8976
3.2398 | -97.7726
3.2435 | -97.6341
3.2472 | -97.5011
3.2508 | -97.393
3.2545 | -97.335
3.2582 | -97.3288
3.2619 | -97.3776
3.2655 | -97.4533
3.2692 | -97.5363
3.2729 | -97.5938
3.2766 | -97.6048
3.2802 | -97.5803
3.2839 | -97.5069
3.2876 | -97.4058
3.2912 | -97.2883
3.2949 | -97.1651
3.2986 | -97.0597
3.3023 | -96.9877
3.3059 | -96.9595
3.3096 | -96.9737
3.3133 | -97.0302
3.317 | -97.1209
3.3206 | -97.2413
3.3243 | -97.3662
3.328 | -97.4859
3.3316 | -97.5827
3.3353 | -97.647
3.339 | -97.6851
3.3427 | -97.6787
3.3463 | -97.6505
3.35 | -97.6006
3.3537 | -97.5436
3.3574 | -97.506
3.361 | -97.5134
3.3647 | -97.5689
3.3684 | -97.6658
3.3721 | -97.7756
3.3757 | -97.8512
3.3794 | -97.8722
3.3831 | -97.8203
3.3867 | -97.7109
3.3904 | -97.5657
3.3941 | -97.4243
3.3978 | -97.3158
3.4014 | -97.2453
3.4051 | -97.2117
3.4088 | -97.2059
3.4125 | -97.212
3.4161 | -97.1974
3.4198 | -97.1656
3.4235 | -97.1229
3.4272 | -97.0714
3.4308 | -97.0343
3.4345 | -97.0287
3.4382 | -97.0634
3.4418 | -97.1505
3.4455 | -97.2896
3.4492 | -97.447
3.4529 | -97.6084
3.4565 | -97.7523
3.4602 | -97.8431
3.4639 | -97.8611
3.4676 | -97.8311
3.4712 | -97.7893
3.4749 | -97.7534
3.4786 | -97.7412
3.4823 | -97.7672
3.4859 | -97.8171
3.4896 | -97.8729
3.4933 | -97.9265
3.4969 | -97.9739
3.5006 | -98.0255
3.5043 | -98.0998
3.508 | -98.1826
3.5116 | -98.2765
3.5153 | -98.3782
3.519 | -98.4681
3.5227 | -98.5283
3.5263 | -98.5479
3.53 | -98.5063
3.5337 | -98.3763
3.5374 | -98.1533
3.541 | -97.8727
3.5447 | -97.5656
3.5484 | -97.2763
3.552 | -97.0384
3.5557 | -96.8838
3.5594 | -96.8339
3.5631 | -96.8787
3.5667 | -97.0007
3.5704 | -97.1803
3.5741 | -97.3779
3.5778 | -97.5574
3.5814 | -97.7014
3.5851 | -97.8067
3.5888 | -97.8895
3.5925 | -97.9571
3.5961 | -98.012
3.5998 | -98.0781
3.6035 | -98.147
3.6071 | -98.2033
3.6108 | -98.227
3.6145 | -98.2155
3.6182 | -98.189
3.6218 | -98.1439
3.6255 | -98.093
3.6292 | -98.0533
3.6329 | -98.0252
3.6365 | -97.995
3.6402 | -97.9647
3.6439 | -97.942
3.6475 | -97.9456
3.6512 | -97.9738
3.6549 | -98.016
3.6586 | -98.0689
3.6622 | -98.1156
3.6659 | -98.1487
3.6696 | -98.1742
3.6733 | -98.1907
3.6769 | -98.2116
3.6806 | -98.2304
3.6843 | -98.2381
3.688 | -98.2315
3.6916 | -98.2096
3.6953 | -98.1641
3.699 | -98.094
3.7026 | -98.0015
3.7063 | -97.9285
3.71 | -97.8779
3.7137 | -97.8485
3.7173 | -97.8493
3.721 | -97.887
3.7247 | -97.9435
3.7284 | -98.0115
3.732 | -98.0802
3.7357 | -98.1364
3.7394 | -98.1617
3.7431 | -98.1591
3.7467 | -98.1261
3.7504 | -98.0805
3.7541 | -98.0281
3.7577 | -97.9819
3.7614 | -97.9573
3.7651 | -97.9482
3.7688 | -97.967
3.7724 | -98.0065
3.7761 | -98.069
3.7798 | -98.144
3.7835 | -98.2262
3.7871 | -98.317
3.7908 | -98.4052
3.7945 | -98.4922
3.7982 | -98.5674
3.8018 | -98.6303
3.8055 | -98.6715
3.8092 | -98.6912
3.8128 | -98.6905
3.8165 | -98.684
3.8202 | -98.6763
3.8239 | -98.6651
3.8275 | -98.6535
3.8312 | -98.6273
3.8349 | -98.5994
3.8386 | -98.5804
3.8422 | -98.5794
3.8459 | -98.6051
3.8496 | -98.6572
3.8533 | -98.7265
3.8569 | -98.7953
3.8606 | -98.8626
3.8643 | -98.9116
3.8679 | -98.9193
3.8716 | -98.8804
3.8753 | -98.7913
3.879 | -98.6651
3.8826 | -98.5145
3.8863 | -98.3648
3.89 | -98.2397
3.8937 | -98.1516
3.8973 | -98.1066
3.901 | -98.1038
3.9047 | -98.1354
3.9084 | -98.1954
3.912 | -98.2629
3.9157 | -98.3251
3.9194 | -98.3813
3.923 | -98.4245
3.9267 | -98.4694
3.9304 | -98.5287
3.9341 | -98.5995
3.9377 | -98.6754
3.9414 | -98.7466
3.9451 | -98.8104
3.9488 | -98.8475
3.9524 | -98.8584
3.9561 | -98.8489
3.9598 | -98.8198
3.9635 | -98.7951
3.9671 | -98.7698
3.9708 | -98.7473
3.9745 | -98.7372
3.9781 | -98.7309
3.9818 | -98.7331
3.9855 | -98.733
3.9892 | -98.7268
3.9928 | -98.7202
3.9965 | -98.7204
4.0002 | -98.7283
4.0039 | -98.7481
4.0075 | -98.7816
4.0112 | -98.8249
4.0149 | -98.8691
4.0186 | -98.9009
4.0222 | -98.9152
4.0259 | -98.9068
4.0296 | -98.9019
4.0332 | -98.907
4.0369 | -98.9327
4.0406 | -98.9802
4.0443 | -99.0446
4.0479 | -99.1293
4.0516 | -99.2203
4.0553 | -99.2809
4.059 | -99.3054
4.0626 | -99.2969
4.0663 | -99.2649
4.07 | -99.2132
4.0736 | -99.1674
4.0773 | -99.1381
4.081 | -99.136
4.0847 | -99.1679
4.0883 | -99.2288
4.092 | -99.3179
4.0957 | -99.4326
4.0994 | -99.5533
4.103 | -99.6698
4.1067 | -99.751
4.1104 | -99.7823
4.1141 | -99.7604
4.1177 | -99.7026
4.1214 | -99.6301
4.1251 | -99.5438
4.1287 | -99.4512
4.1324 | -99.3753
4.1361 | -99.3059
4.1398 | -99.2509
4.1434 | -99.2105
4.1471 | -99.1762
4.1508 | -99.1477
4.1545 | -99.1252
4.1581 | -99.1056
4.1618 | -99.0983
4.1655 | -99.0968
4.1692 | -99.1162
4.1728 | -99.1531
4.1765 | -99.1957
4.1802 | -99.2384
4.1838 | -99.2694
4.1875 | -99.298
4.1912 | -99.3344
4.1949 | -99.3776
4.1985 | -99.4304
4.2022 | -99.4874
4.2059 | -99.5465
4.2096 | -99.5928
4.2132 | -99.6247
4.2169 | -99.6341
4.2206 | -99.6226
4.2243 | -99.5916
4.2279 | -99.5506
4.2316 | -99.5056
4.2353 | -99.4618
4.2389 | -99.4529
4.2426 | -99.4397
4.2463 | -99.4365
4.25 | -99.4395
4.2536 | -99.4498
4.2573 | -99.4609
4.261 | -99.4622
4.2647 | -99.466
4.2683 | -99.4747
4.272 | -99.4866
4.2757 | -99.5105
4.2794 | -99.5438
4.283 | -99.5904
4.2867 | -99.6484
4.2904 | -99.7177
4.294 | -99.7982
4.2977 | -99.8887
4.3014 | -99.9907
4.3051 | -100.0899
4.3087 | -100.1878
4.3124 | -100.278
4.3161 | -100.3612
4.3198 | -100.4239
4.3234 | -100.4631
4.3271 | -100.4764
4.3308 | -100.482
4.3345 | -100.4679
4.3381 | -100.424
4.3418 | -100.3623
4.3455 | -100.2834
4.3491 | -100.1989
4.3528 | -100.1026
4.3565 | -99.9982
4.3602 | -99.8991
4.3638 | -99.8065
4.3675 | -99.7192
4.3712 | -99.6452
4.3749 | -99.5882
4.3785 | -99.5446
4.3822 | -99.513
4.3859 | -99.4949
4.3895 | -99.4958
4.3932 | -99.5106
4.3969 | -99.5277
4.4006 | -99.5384
4.4042 | -99.5461
4.4079 | -99.5602
4.4116 | -99.5837
4.4153 | -99.6174
4.4189 | -99.6589
4.4226 | -99.723
4.4263 | -99.806
4.43 | -99.8881
4.4336 | -99.9661
4.4373 | -100.028
4.441 | -100.0673
4.4446 | -100.074
4.4483 | -100.0548
4.452 | -100.0167
4.4557 | -99.9672
4.4593 | -99.8943
4.463 | -99.807
4.4667 | -99.7409
4.4704 | -99.6903
4.474 | -99.6486
4.4777 | -99.6252
4.4814 | -99.6203
4.4851 | -99.6273
4.4887 | -99.6362
4.4924 | -99.6482
4.4961 | -99.6581
4.4997 | -99.6716
4.5034 | -99.7018
4.5071 | -99.7536
4.5108 | -99.8228
4.5144 | -99.9095
4.5181 | -99.9894
4.5218 | -100.0568
4.5255 | -100.113
4.5291 | -100.1652
4.5328 | -100.2027
4.5365 | -100.2007
4.5402 | -100.1896
4.5438 | -100.1781
4.5475 | -100.1442
4.5512 | -100.096
4.5548 | -100.0532
4.5585 | -100.0036
4.5622 | -99.9454
4.5659 | -99.8942
4.5695 | -99.8495
4.5732 | -99.8072
4.5769 | -99.776
4.5806 | -99.7626
4.5842 | -99.7618
4.5879 | -99.7743
4.5916 | -99.7976
4.5953 | -99.8202
4.5989 | -99.8457
4.6026 | -99.8757
4.6063 | -99.9179
4.6099 | -99.9647
4.6136 | -100.0194
4.6173 | -100.0915
4.621 | -100.1731
4.6246 | -100.258
4.6283 | -100.3561
4.632 | -100.4487
4.6357 | -100.5189
4.6393 | -100.5776
4.643 | -100.6212
4.6467 | -100.6376
4.6504 | -100.6204
4.654 | -100.5855
4.6577 | -100.5393
4.6614 | -100.485
4.665 | -100.4361
4.6687 | -100.3987
4.6724 | -100.3745
4.6761 | -100.3643
4.6797 | -100.3645
4.6834 | -100.359
4.6871 | -100.3436
4.6908 | -100.3263
4.6944 | -100.2995
4.6981 | -100.2585
4.7018 | -100.2108
4.7054 | -100.161
4.7091 | -100.0962
4.7128 | -100.0319
4.7165 | -99.981
4.7201 | -99.9514
4.7238 | -99.9366
4.7275 | -99.9393
4.7312 | -99.9522
4.7348 | -99.9746
4.7385 | -100.0081
4.7422 | -100.0588
4.7459 | -100.1171
4.7495 | -100.1853
4.7532 | -100.2696
4.7569 | -100.3572
4.7605 | -100.4349
4.7642 | -100.4965
4.7679 | -100.5134
4.7716 | -100.5402
4.7752 | -100.5409
4.7789 | -100.524
4.7826 | -100.4928
4.7863 | -100.4503
4.7899 | -100.4029
4.7936 | -100.353
4.7973 | -100.3108
4.801 | -100.2841
4.8046 | -100.272
4.8083 | -100.2716
4.812 | -100.2832
4.8156 | -100.3035
4.8193 | -100.3365
4.823 | -100.3752
4.8267 | -100.4113
4.8303 | -100.4514
4.834 | -100.4944
4.8377 | -100.5375
4.8414 | -100.5743
4.845 | -100.617
4.8487 | -100.6658
4.8524 | -100.7134
4.8561 | -100.7481
4.8597 | -100.7678
4.8634 | -100.7587
4.8671 | -100.7203
4.8707 | -100.6612
4.8744 | -100.584
4.8781 | -100.4972
4.8818 | -100.4187
4.8854 | -100.3564
4.8891 | -100.3137
4.8928 | -100.2919
4.8965 | -100.3003
4.9001 | -100.3356
4.9038 | -100.3993
4.9075 | -100.4853
4.9112 | -100.588
4.9148 | -100.7086
4.9185 | -100.8369
4.9222 | -100.9483
4.9258 | -101.0438
4.9295 | -101.1186
4.9332 | -101.152
4.9369 | -101.1392
4.9405 | -101.0836
4.9442 | -100.9889
4.9479 | -100.868
4.9516 | -100.7466
4.9552 | -100.6295
4.9589 | -100.523
4.9626 | -100.4289
4.9663 | -100.3597
4.9699 | -100.3202
4.9736 | -100.2941
4.9773 | -100.2726
4.9809 | -100.2537
4.9846 | -100.2435
4.9883 | -100.238
4.992 | -100.2316
4.9956 | -100.2272
4.9993 | -100.2371
5.003 | -100.256
5.0067 | -100.2887
5.0103 | -100.345
5.014 | -100.4288
5.0177 | -100.5326
5.0214 | -100.6523
5.025 | -100.7832
5.0287 | -100.9131
5.0324 | -101.0369
5.036 | -101.1577
5.0397 | -101.2676
5.0434 | -101.3471
5.0471 | -101.3991
5.0507 | -101.4283
5.0544 | -101.4408
5.0581 | -101.436
5.0618 | -101.4115
5.0654 | -101.378
5.0691 | -101.3452
5.0728 | -101.3126
5.0765 | -101.2743
5.0801 | -101.233
5.0838 | -101.1863
5.0875 | -101.1334
5.0911 | -101.0717
5.0948 | -101.0121
5.0985 | -100.9524
5.1022 | -100.8934
5.1058 | -100.8467
5.1095 | -100.8152
5.1132 | -100.7908
5.1169 | -100.7829
5.1205 | -100.7911
5.1242 | -100.8162
5.1279 | -100.861
5.1315 | -100.9064
5.1352 | -100.9436
5.1389 | -100.9658
5.1426 | -100.9759
5.1462 | -100.9642
5.1499 | -100.9388
5.1536 | -100.9099
5.1573 | -100.8851
5.1609 | -100.8587
5.1646 | -100.845
5.1683 | -100.8483
5.172 | -100.8682
5.1756 | -100.9101
5.1793 | -100.9643
5.183 | -101.0237
5.1866 | -101.086
5.1903 | -101.1515
5.194 | -101.2071
5.1977 | -101.2541
5.2013 | -101.2801
5.205 | -101.2803
5.2087 | -101.2724
5.2124 | -101.2547
5.216 | -101.2334
5.2197 | -101.2121
5.2234 | -101.1962
5.2271 | -101.1822
5.2307 | -101.1805
5.2344 | -101.194
5.2381 | -101.2281
5.2417 | -101.2748
5.2454 | -101.3234
5.2491 | -101.3679
5.2528 | -101.4138
5.2564 | -101.4507
5.2601 | -101.4743
5.2638 | -101.4794
5.2675 | -101.4593
5.2711 | -101.4214
5.2748 | -101.3685
5.2785 | -101.3093
5.2822 | -101.244
5.2858 | -101.1702
5.2895 | -101.0985
5.2932 | -101.0348
5.2968 | -101.0051
5.3005 | -100.962
5.3042 | -100.9325
5.3079 | -100.9119
5.3115 | -100.902
5.3152 | -100.9006
5.3189 | -100.906
5.3226 | -100.9134
5.3262 | -100.9211
5.3299 | -100.9301
5.3336 | -100.9291
5.3373 | -100.9188
5.3409 | -100.9029
5.3446 | -100.8858
5.3483 | -100.8612
5.3519 | -100.8363
5.3556 | -100.8114
5.3593 | -100.7877
5.363 | -100.7675
5.3666 | -100.7507
5.3703 | -100.7403
5.374 | -100.7392
5.3777 | -100.7491
5.3813 | -100.7706
5.385 | -100.8046
5.3887 | -100.8471
5.3924 | -100.8992
5.396 | -100.9477
5.3997 | -100.9945
5.4034 | -101.0385
5.407 | -101.0816
5.4107 | -101.1253
5.4144 | -101.1664
5.4181 | -101.2061
5.4217 | -101.252
5.4254 | -101.2942
5.4291 | -101.3264
5.4328 | -101.3552
5.4364 | -101.3723
5.4401 | -101.3709
5.4438 | -101.3622
5.4474 | -101.3405
5.4511 | -101.3115
5.4548 | -101.2777
5.4585 | -101.2487
5.4621 | -101.2238
5.4658 | -101.206
5.4695 | -101.1958
5.4732 | -101.19
5.4768 | -101.1961
5.4805 | -101.223
5.4842 | -101.2639
5.4879 | -101.3173
5.4915 | -101.3756
5.4952 | -101.4394
5.4989 | -101.4984
5.5025 | -101.5511
5.5062 | -101.5878
5.5099 | -101.6043
5.5136 | -101.6028
5.5172 | -101.5812
5.5209 | -101.5385
5.5246 | -101.4907
5.5283 | -101.4351
5.5319 | -101.3885
5.5356 | -101.3437
5.5393 | -101.316
5.543 | -101.3054
5.5466 | -101.3025
5.5503 | -101.3125
5.554 | -101.3281
5.5576 | -101.3478
5.5613 | -101.3667
5.565 | -101.3822
5.5687 | -101.3932
5.5723 | -101.3957
5.576 | -101.3967
5.5797 | -101.3862
5.5834 | -101.3772
5.587 | -101.3665
5.5907 | -101.3507
5.5944 | -101.34
5.5981 | -101.3299
5.6017 | -101.3217
5.6054 | -101.3279
5.6091 | -101.3395
5.6127 | -101.3579
5.6164 | -101.3844
5.6201 | -101.422
5.6238 | -101.4599
5.6274 | -101.4973
5.6311 | -101.5489
5.6348 | -101.5959
5.6385 | -101.6404
5.6421 | -101.6782
5.6458 | -101.7051
5.6495 | -101.7269
5.6532 | -101.7457
5.6568 | -101.7587
5.6605 | -101.7597
5.6642 | -101.7544
5.6678 | -101.7368
5.6715 | -101.7147
5.6752 | -101.6961
5.6789 | -101.6832
5.6825 | -101.6845
5.6862 | -101.6932
5.6899 | -101.705
5.6936 | -101.7264
5.6972 | -101.7484
5.7009 | -101.7581
5.7046 | -101.7617
5.7083 | -101.7475
5.7119 | -101.7161
5.7156 | -101.6592
5.7193 | -101.59
5.7229 | -101.5159
5.7266 | -101.4457
5.7303 | -101.3896
5.734 | -101.3454
5.7376 | -101.3163
5.7413 | -101.3099
5.745 | -101.3229
5.7487 | -101.3542
5.7523 | -101.4048
5.756 | -101.468
5.7597 | -101.5476
5.7633 | -101.6225
5.767 | -101.6974
5.7707 | -101.7713
5.7744 | -101.8254
5.778 | -101.8563
5.7817 | -101.8646
5.7854 | -101.8559
5.7891 | -101.82
5.7927 | -101.7694
5.7964 | -101.7105
5.8001 | -101.6463
5.8038 | -101.5781
5.8074 | -101.5098
5.8111 | -101.4393
5.8148 | -101.3727
5.8184 | -101.3205
5.8221 | -101.2887
5.8258 | -101.3215
5.8295 | -101.334
5.8331 | -101.3651
5.8368 | -101.4174
5.8405 | -101.4878
5.8442 | -101.5789
5.8478 | -101.6664
5.8515 | -101.756
5.8552 | -101.8364
5.8589 | -101.9031
5.8625 | -101.9542
5.8662 | -101.9904
5.8699 | -102.0029
5.8735 | -101.9915
5.8772 | -101.9774
5.8809 | -101.9534
5.8846 | -101.9258
5.8882 | -101.8998
5.8919 | -101.8677
5.8956 | -101.8387
5.8993 | -101.818
5.9029 | -101.7989
5.9066 | -101.7781
5.9103 | -101.7523
5.914 | -101.7301
5.9176 | -101.7105
5.9213 | -101.693
5.925 | -101.6782
5.9286 | -101.6641
5.9323 | -101.6511
5.936 | -101.6342
5.9397 | -101.6307
5.9433 | -101.6344
5.947 | -101.646
5.9507 | -101.672
5.9544 | -101.7108
5.958 | -101.7571
5.9617 | -101.8148
5.9654 | -101.8779
5.9691 | -101.9301
5.9727 | -101.9768
5.9764 | -102.0125
5.9801 | -102.0315
5.9837 | -102.0384
5.9874 | -102.0316
5.9911 | -102.0143
5.9948 | -101.9932
5.9984 | -101.9692
6.0021 | -101.941
6.0058 | -101.9088
6.0095 | -101.8759
6.0131 | -101.8406
6.0168 | -101.8112
6.0205 | -101.7914
6.0242 | -101.7772
6.0278 | -101.7728
6.0315 | -101.7689
6.0352 | -101.7719
6.0388 | -101.77
6.0425 | -101.7618
6.0462 | -101.7559
6.0499 | -101.7503
6.0535 | -101.7423
6.0572 | -101.7386
6.0609 | -101.7422
6.0646 | -101.7533
6.0682 | -101.7679
6.0719 | -101.7826
6.0756 | -101.8017
6.0793 | -101.8187
6.0829 | -101.8358
6.0866 | -101.8478
6.0903 | -101.85
6.0939 | -101.8522
6.0976 | -101.8477
6.1013 | -101.8355
6.105 | -101.8253
6.1086 | -101.8182
6.1123 | -101.8151
6.116 | -101.8223
6.1197 | -101.8392
6.1233 | -101.8699
6.127 | -101.9084
6.1307 | -101.9547
6.1344 | -102.0114
6.138 | -102.0668
6.1417 | -102.1239
6.1454 | -102.1808
6.149 | -102.2313
6.1527 | -102.2761
6.1564 | -102.3089
6.1601 | -102.3331
6.1637 | -102.3538
6.1674 | -102.3579
6.1711 | -102.356
6.1748 | -102.3508
6.1784 | -102.343
6.1821 | -102.3415
6.1858 | -102.3367
6.1894 | -102.3264
6.1931 | -102.3211
6.1968 | -102.3169
6.2005 | -102.2973
6.2041 | -102.2738
6.2078 | -102.2523
6.2115 | -102.2301
6.2152 | -102.2012
6.2188 | -102.1702
6.2225 | -102.1475
6.2262 | -102.1231
6.2299 | -102.1047
6.2335 | -102.0912
6.2372 | -102.0849
6.2409 | -102.0803
6.2445 | -102.08
6.2482 | -102.0862
6.2519 | -102.0913
6.2556 | -102.0957
6.2592 | -102.0977
6.2629 | -102.092
6.2666 | -102.0836
6.2703 | -102.0812
6.2739 | -102.0785
6.2776 | -102.0836
6.2813 | -102.1037
6.285 | -102.1336
6.2886 | -102.1751
6.2923 | -102.2301
6.296 | -102.2892
6.2996 | -102.3523
6.3033 | -102.4137
6.307 | -102.4717
6.3107 | -102.5146
6.3143 | -102.535
6.318 | -102.5408
6.3217 | -102.5242
6.3254 | -102.4836
6.329 | -102.43
6.3327 | -102.3646
6.3364 | -102.2867
6.3401 | -102.1982
6.3437 | -102.1107
6.3474 | -102.0235
6.3511 | -101.949
6.3547 | -101.8836
6.3584 | -101.8612
6.3621 | -101.8268
6.3658 | -101.8095
6.3694 | -101.8099
6.3731 | -101.823
6.3768 | -101.8473
6.3805 | -101.886
6.3841 | -101.9362
6.3878 | -101.993
6.3915 | -102.054
6.3952 | -102.1186
6.3988 | -102.1882
6.4025 | -102.2612
6.4062 | -102.3383
6.4098 | -102.414
6.4135 | -102.4848
6.4172 | -102.5565
6.4209 | -102.6257
6.4245 | -102.684
6.4282 | -102.7378
6.4319 | -102.7852
6.4356 | -102.8247
6.4392 | -102.8495
6.4429 | -102.8628
6.4466 | -102.8626
6.4502 | -102.8547
6.4539 | -102.8422
6.4576 | -102.8208
6.4613 | -102.7886
6.4649 | -102.7569
6.4686 | -102.7275
6.4723 | -102.704
6.476 | -102.6881
6.4796 | -102.6814
6.4833 | -102.6859
6.487 | -102.7022
6.4907 | -102.7235
6.4943 | -102.7436
6.498 | -102.7712
6.5017 | -102.8015
6.5053 | -102.8288
6.509 | -102.8518
6.5127 | -102.8604
6.5164 | -102.8614
6.52 | -102.8581
6.5237 | -102.8438
6.5274 | -102.8183
6.5311 | -102.7814
6.5347 | -102.7263
6.5384 | -102.6627
6.5421 | -102.5962
6.5458 | -102.5321
6.5494 | -102.4729
6.5531 | -102.4173
6.5568 | -102.3738
6.5604 | -102.3485
6.5641 | -102.3313
6.5678 | -102.3329
6.5715 | -102.3486
6.5751 | -102.372
6.5788 | -102.4017
6.5825 | -102.4394
6.5862 | -102.4816
6.5898 | -102.521
6.5935 | -102.557
6.5972 | -102.5748
6.6009 | -102.585
6.6045 | -102.5928
6.6082 | -102.5945
6.6119 | -102.5923
6.6155 | -102.5825
6.6192 | -102.5653
6.6229 | -102.5548
6.6266 | -102.5549
6.6302 | -102.5663
6.6339 | -102.5916
6.6376 | -102.619
6.6413 | -102.6502
6.6449 | -102.6835
6.6486 | -102.7212
6.6523 | -102.7672
6.656 | -102.8103
6.6596 | -102.8451
6.6633 | -102.8751
6.667 | -102.8885
6.6706 | -102.8828
6.6743 | -102.8758
6.678 | -102.858
6.6817 | -102.8355
6.6853 | -102.8176
6.689 | -102.8016
6.6927 | -102.7894
6.6964 | -102.781
6.7 | -102.7731
6.7037 | -102.7666
6.7074 | -102.7654
6.7111 | -102.7657
6.7147 | -102.7676
6.7184 | -102.7772
6.7221 | -102.7832
6.7257 | -102.7907
6.7294 | -102.8052
6.7331 | -102.8313
6.7368 | -102.8656
6.7404 | -102.9016
6.7441 | -102.9418
6.7478 | -102.9805
6.7515 | -103.0187
6.7551 | -103.0556
6.7588 | -103.0863
6.7625 | -103.1035
6.7662 | -103.1134
6.7698 | -103.112
6.7735 | -103.0861
6.7772 | -103.0502
6.7808 | -103.0045
6.7845 | -102.9529
6.7882 | -102.9046
6.7919 | -102.8613
6.7955 | -102.8266
6.7992 | -102.7982
6.8029 | -102.7806
6.8066 | -102.7712
6.8102 | -102.7708
6.8139 | -102.7712
6.8176 | -102.7757
6.8213 | -102.7864
6.8249 | -102.7951
6.8286 | -102.7979
6.8323 | -102.8017
6.8359 | -102.8144
6.8396 | -102.828
6.8433 | -102.8336
6.847 | -102.8402
6.8506 | -102.8447
6.8543 | -102.8497
6.858 | -102.8562
6.8617 | -102.8608
6.8653 | -102.8623
6.869 | -102.8668
6.8727 | -102.8786
6.8763 | -102.8938
6.88 | -102.9116
6.8837 | -102.9338
6.8874 | -102.9565
6.891 | -102.9777
6.8947 | -102.9948
6.8984 | -103.0097
6.9021 | -103.021
6.9057 | -103.0217
6.9094 | -103.0254
6.9131 | -103.0297
6.9168 | -103.0343
6.9204 | -103.0354
6.9241 | -103.039
6.9278 | -103.0422
6.9314 | -103.0533
6.9351 | -103.0648
6.9388 | -103.0729
6.9425 | -103.0817
6.9461 | -103.0936
6.9498 | -103.097
6.9535 | -103.0956
6.9572 | -103.093
6.9608 | -103.0861
6.9645 | -103.08
6.9682 | -103.0697
6.9719 | -103.0552
6.9755 | -103.0397
6.9792 | -103.0295
6.9829 | -103.0244
6.9865 | -103.0231
6.9902 | -103.0219
6.9939 | -103.0285
6.9976 | -103.0376
7.0012 | -103.0548
7.0049 | -103.0763
7.0086 | -103.0942
7.0123 | -103.1185
7.0159 | -103.1321
7.0196 | -103.1402
7.0233 | -103.1436
7.027 | -103.1372
7.0306 | -103.1224
7.0343 | -103.1034
7.038 | -103.0717
7.0416 | -103.0336
7.0453 | -102.993
7.049 | -102.9561
7.0527 | -102.9232
7.0563 | -102.897
7.06 | -102.8744
7.0637 | -102.8578
7.0674 | -102.8487
7.071 | -102.8473
7.0747 | -102.8539
7.0784 | -102.8651
7.0821 | -102.8746
7.0857 | -102.881
7.0894 | -102.8907
7.0931 | -102.9045
7.0967 | -102.9206
7.1004 | -102.9334
7.1041 | -102.9476
7.1078 | -102.9579
7.1114 | -102.9652
7.1151 | -102.9758
7.1188 | -102.9855
7.1225 | -102.9949
7.1261 | -103.0002
7.1298 | -103.0112
7.1335 | -103.0257
7.1372 | -103.0381
7.1408 | -103.0475
7.1445 | -103.0583
7.1482 | -103.0672
7.1518 | -103.071
7.1555 | -103.0695
7.1592 | -103.0601
7.1629 | -103.0437
7.1665 | -103.0261
7.1702 | -103.0019
7.1739 | -102.9737
7.1776 | -102.9448
7.1812 | -102.913
7.1849 | -102.8841
7.1886 | -102.8603
7.1922 | -102.8441
7.1959 | -102.8333
7.1996 | -102.8269
7.2033 | -102.8257
7.2069 | -102.8347
7.2106 | -102.8514
7.2143 | -102.8774
7.218 | -102.9073
7.2216 | -102.9354
7.2253 | -102.9698
7.229 | -103.0059
7.2327 | -103.044
7.2363 | -103.0663
7.24 | -103.0796
7.2437 | -103.0897
7.2473 | -103.0889
7.251 | -103.0767
7.2547 | -103.0513
7.2584 | -103.0229
7.262 | -102.9932
7.2657 | -102.9624
7.2694 | -102.9355
7.2731 | -102.9126
7.2767 | -102.8987
7.2804 | -102.8939
7.2841 | -102.8956
7.2878 | -102.9074
7.2914 | -102.9191
7.2951 | -102.9319
7.2988 | -102.9484
7.3024 | -102.9632
7.3061 | -102.9724
7.3098 | -102.9858
7.3135 | -102.9951
7.3171 | -102.9987
7.3208 | -103.0029
7.3245 | -103.0039
7.3282 | -103.0096
7.3318 | -103.009
7.3355 | -103.0111
7.3392 | -103.0122
7.3429 | -103.0199
7.3465 | -103.0301
7.3502 | -103.0431
7.3539 | -103.065
7.3575 | -103.093
7.3612 | -103.1269
7.3649 | -103.1623
7.3686 | -103.1967
7.3722 | -103.2344
7.3759 | -103.2664
7.3796 | -103.2949
7.3833 | -103.3248
7.3869 | -103.3477
7.3906 | -103.3643
7.3943 | -103.3757
7.398 | -103.3888
7.4016 | -103.4007
7.4053 | -103.4063
7.409 | -103.4116
7.4126 | -103.4225
7.4163 | -103.4357
7.42 | -103.4443
7.4237 | -103.4535
7.4273 | -103.4621
7.431 | -103.4681
7.4347 | -103.4788
7.4384 | -103.4836
7.442 | -103.4842
7.4457 | -103.4833
7.4494 | -103.4814
7.4531 | -103.4758
7.4567 | -103.4676
7.4604 | -103.4573
7.4641 | -103.4462
7.4677 | -103.4393
7.4714 | -103.4382
7.4751 | -103.4353
7.4788 | -103.4356
7.4824 | -103.437
7.4861 | -103.4434
7.4898 | -103.4506
7.4935 | -103.4581
7.4971 | -103.4654
7.5008 | -103.4754
7.5045 | -103.4863
7.5081 | -103.4945
7.5118 | -103.4983
7.5155 | -103.4977
7.5192 | -103.4967
7.5228 | -103.491
7.5265 | -103.482
7.5302 | -103.4708
7.5339 | -103.455
7.5375 | -103.4375
7.5412 | -103.424
7.5449 | -103.4105
7.5486 | -103.3933
7.5522 | -103.3804
7.5559 | -103.3759
7.5596 | -103.3723
7.5632 | -103.3739
7.5669 | -103.3813
7.5706 | -103.3933
7.5743 | -103.4096
7.5779 | -103.4256
7.5816 | -103.4392
7.5853 | -103.4543
7.589 | -103.4716
7.5926 | -103.4824
7.5963 | -103.4929
7.6 | -103.4988
7.6037 | -103.4976
7.6073 | -103.4951
7.611 | -103.4916
7.6147 | -103.486
7.6183 | -103.4764
7.622 | -103.4715
7.6257 | -103.4672
7.6294 | -103.4595
7.633 | -103.4487
7.6367 | -103.4368
7.6404 | -103.4197
7.6441 | -103.4025
7.6477 | -103.384
7.6514 | -103.3667
7.6551 | -103.3485
7.6588 | -103.3297
7.6624 | -103.3166
7.6661 | -103.305
7.6698 | -103.2987
7.6734 | -103.2978
7.6771 | -103.3006
7.6808 | -103.3042
7.6845 | -103.3087
7.6881 | -103.3133
7.6918 | -103.323
7.6955 | -103.3337
7.6992 | -103.3396
7.7028 | -103.3456
7.7065 | -103.3549
7.7102 | -103.3639
7.7139 | -103.3755
7.7175 | -103.3898
7.7212 | -103.4041
7.7249 | -103.4202
7.7285 | -103.4371
7.7322 | -103.4601
7.7359 | -103.4775
7.7396 | -103.4984
7.7432 | -103.524
7.7469 | -103.549
7.7506 | -103.5728
7.7543 | -103.6017
7.7579 | -103.6258
7.7616 | -103.6454
7.7653 | -103.6582
7.769 | -103.6702
7.7726 | -103.6874
7.7763 | -103.7026
7.78 | -103.715
7.7836 | -103.7225
7.7873 | -103.7332
7.791 | -103.7424
7.7947 | -103.7514
7.7983 | -103.7602
7.802 | -103.7695
7.8057 | -103.7739
7.8094 | -103.7713
7.813 | -103.7707
7.8167 | -103.7699
7.8204 | -103.7635
7.8241 | -103.7511
7.8277 | -103.7397
7.8314 | -103.7325
7.8351 | -103.7248
7.8387 | -103.7137
7.8424 | -103.7017
7.8461 | -103.6925
7.8498 | -103.6878
7.8534 | -103.6792
7.8571 | -103.669
7.8608 | -103.6593
7.8645 | -103.6533
7.8681 | -103.6498
7.8718 | -103.6551
7.8755 | -103.6611
7.8792 | -103.6697
7.8828 | -103.6802
7.8865 | -103.6909
7.8902 | -103.7055
7.8938 | -103.7205
7.8975 | -103.7396
7.9012 | -103.76
7.9049 | -103.7833
7.9085 | -103.8078
7.9122 | -103.8293
7.9159 | -103.8523
7.9196 | -103.8779
7.9232 | -103.9105
7.9269 | -103.9456
7.9306 | -103.981
7.9342 | -104.0204
7.9379 | -104.0649
7.9416 | -104.1065
7.9453 | -104.1459
7.9489 | -104.1618
7.9526 | -104.1819
7.9563 | -104.2061
7.96 | -104.217
7.9636 | -104.2161
7.9673 | -104.2068
7.971 | -104.1906
7.9747 | -104.1654
7.9783 | -104.1319
7.982 | -104.0901
7.9857 | -104.0406
7.9893 | -103.9864
7.993 | -103.9382
7.9967 | -103.8913
8.0004 | -103.8458
8.004 | -103.8044
8.0077 | -103.7675
8.0114 | -103.737
8.0151 | -103.7093
8.0187 | -103.6848
8.0224 | -103.6657
8.0261 | -103.6527
8.0298 | -103.6415
8.0334 | -103.6296
8.0371 | -103.6194
8.0408 | -103.6036
8.0444 | -103.5878
8.0481 | -103.5711
8.0518 | -103.5551
8.0555 | -103.5385
8.0591 | -103.523
8.0628 | -103.5054
8.0665 | -103.4923
8.0702 | -103.4825
8.0738 | -103.4696
8.0775 | -103.4563
8.0812 | -103.4497
8.0849 | -103.4473
8.0885 | -103.4488
8.0922 | -103.4582
8.0959 | -103.4732
8.0995 | -103.4964
8.1032 | -103.5227
8.1069 | -103.5505
8.1106 | -103.5846
8.1142 | -103.6192
8.1179 | -103.6538
8.1216 | -103.6859
8.1253 | -103.7143
8.1289 | -103.7373
8.1326 | -103.761
8.1363 | -103.7831
8.14 | -103.8005
8.1436 | -103.8128
8.1473 | -103.8245
8.151 | -103.8332
8.1546 | -103.833
8.1583 | -103.8316
8.162 | -103.832
8.1657 | -103.8222
8.1693 | -103.809
8.173 | -103.7936
8.1767 | -103.7784
8.1804 | -103.765
8.184 | -103.7559
8.1877 | -103.7451
8.1914 | -103.7363
8.1951 | -103.7325
8.1987 | -103.732
8.2024 | -103.7352
8.2061 | -103.7398
8.2097 | -103.7494
8.2134 | -103.7558
8.2171 | -103.7575
8.2208 | -103.7619
8.2244 | -103.7709
8.2281 | -103.7791
8.2318 | -103.7809
8.2355 | -103.7806
8.2391 | -103.7823
8.2428 | -103.7908
8.2465 | -103.8005
8.2501 | -103.8089
8.2538 | -103.8167
8.2575 | -103.8269
8.2612 | -103.8423
8.2648 | -103.8539
8.2685 | -103.8636
8.2722 | -103.8818
8.2759 | -103.9062
8.2795 | -103.927
8.2832 | -103.9485
8.2869 | -103.9742
8.2906 | -104.0001
8.2942 | -104.0225
8.2979 | -104.0367
8.3016 | -104.0482
8.3052 | -104.0565
8.3089 | -104.0636
8.3126 | -104.067
8.3163 | -104.065
8.3199 | -104.0558
8.3236 | -104.0459
8.3273 | -104.0357
8.331 | -104.0249
8.3346 | -104.0171
8.3383 | -104.0148
8.342 | -104.0134
8.3457 | -104.0152
8.3493 | -104.0243
8.353 | -104.037
8.3567 | -104.0523
8.3603 | -104.0684
8.364 | -104.0796
8.3677 | -104.0943
8.3714 | -104.1114
8.375 | -104.1223
8.3787 | -104.1271
8.3824 | -104.1327
8.3861 | -104.1453
8.3897 | -104.1561
8.3934 | -104.1628
8.3971 | -104.1705
8.4008 | -104.18
8.4044 | -104.1878
8.4081 | -104.1968
8.4118 | -104.209
8.4154 | -104.2224
8.4191 | -104.2428
8.4228 | -104.2641
8.4265 | -104.283
8.4301 | -104.2999
8.4338 | -104.3242
8.4375 | -104.3472
8.4412 | -104.3733
8.4448 | -104.3985
8.4485 | -104.4245
8.4522 | -104.449
8.4559 | -104.4759
8.4595 | -104.504
8.4632 | -104.5241
8.4669 | -104.5421
8.4705 | -104.5577
8.4742 | -104.5735
8.4779 | -104.5727
8.4816 | -104.576
8.4852 | -104.5751
8.4889 | -104.5703
8.4926 | -104.5578
8.4963 | -104.5428
8.4999 | -104.5302
8.5036 | -104.5165
8.5073 | -104.4944
8.511 | -104.4678
8.5146 | -104.4404
8.5183 | -104.4141
8.522 | -104.3894
8.5256 | -104.365
8.5293 | -104.3406
8.533 | -104.3149
8.5367 | -104.2863
8.5403 | -104.2554
8.544 | -104.2241
8.5477 | -104.191
8.5514 | -104.1625
8.555 | -104.1394
8.5587 | -104.1143
8.5624 | -104.0904
8.5661 | -104.0741
8.5697 | -104.0641
8.5734 | -104.0563
8.5771 | -104.0551
8.5807 | -104.0578
8.5844 | -104.062
8.5881 | -104.0689
8.5918 | -104.0822
8.5954 | -104.1028
8.5991 | -104.1238
8.6028 | -104.1437
8.6065 | -104.1688
8.6101 | -104.1922
8.6138 | -104.2166
8.6175 | -104.2429
8.6211 | -104.2694
8.6248 | -104.2907
8.6285 | -104.3048
8.6322 | -104.3141
8.6358 | -104.3212
8.6395 | -104.3248
8.6432 | -104.3238
8.6469 | -104.3212
8.6505 | -104.3163
8.6542 | -104.3077
8.6579 | -104.2984
8.6616 | -104.2872
8.6652 | -104.2761
8.6689 | -104.2678
8.6726 | -104.2585
8.6762 | -104.2527
8.6799 | -104.2472
8.6836 | -104.2408
8.6873 | -104.2347
8.6909 | -104.2352
8.6946 | -104.2349
8.6983 | -104.2351
8.702 | -104.2412
8.7056 | -104.2457
8.7093 | -104.2478
8.713 | -104.2571
8.7167 | -104.2693
8.7203 | -104.2774
8.724 | -104.2766
8.7277 | -104.2756
8.7313 | -104.272
8.735 | -104.2651
8.7387 | -104.2588
8.7424 | -104.2469
8.746 | -104.2305
8.7497 | -104.2137
8.7534 | -104.1968
8.7571 | -104.1803
8.7607 | -104.1663
8.7644 | -104.1513
8.7681 | -104.1367
8.7718 | -104.1247
8.7754 | -104.1145
8.7791 | -104.1135
8.7828 | -104.1238
8.7864 | -104.1402
8.7901 | -104.16
8.7938 | -104.1802
8.7975 | -104.2037
8.8011 | -104.2303
8.8048 | -104.254
8.8085 | -104.2757
8.8122 | -104.2975
8.8158 | -104.3133
8.8195 | -104.3221
8.8232 | -104.3286
8.8269 | -104.3307
8.8305 | -104.3263
8.8342 | -104.3143
8.8379 | -104.301
8.8415 | -104.2759
8.8452 | -104.247
8.8489 | -104.2166
8.8526 | -104.1885
8.8562 | -104.1606
8.8599 | -104.1358
8.8636 | -104.1121
8.8673 | -104.089
8.8709 | -104.067
8.8746 | -104.0521
8.8783 | -104.0415
8.882 | -104.0388
8.8856 | -104.0454
8.8893 | -104.0565
8.893 | -104.0753
8.8966 | -104.0995
8.9003 | -104.1306
8.904 | -104.1675
8.9077 | -104.2065
8.9113 | -104.2463
8.915 | -104.2887
8.9187 | -104.3368
8.9224 | -104.3903
8.926 | -104.442
8.9297 | -104.489
8.9334 | -104.5313
8.937 | -104.5746
8.9407 | -104.6118
8.9444 | -104.6476
8.9481 | -104.6761
8.9517 | -104.6961
8.9554 | -104.7136
8.9591 | -104.729
8.9628 | -104.7454
8.9664 | -104.7533
8.9701 | -104.7636
8.9738 | -104.768
8.9775 | -104.7705
8.9811 | -104.7717
8.9848 | -104.7727
8.9885 | -104.7713
8.9921 | -104.769
8.9958 | -104.7678
8.9995 | -104.7723
9.0032 | -104.7772
9.0068 | -104.7668
9.0105 | -104.7588
9.0142 | -104.7448
9.0179 | -104.7231
9.0215 | -104.7019
9.0252 | -104.6761
9.0289 | -104.6432
9.0326 | -104.6024
9.0362 | -104.5608
9.0399 | -104.5152
9.0436 | -104.4685
9.0472 | -104.4225
9.0509 | -104.3703
9.0546 | -104.3188
9.0583 | -104.2708
9.0619 | -104.2318
9.0656 | -104.2029
9.0693 | -104.1751
9.073 | -104.1546
9.0766 | -104.1416
9.0803 | -104.1354
9.084 | -104.1346
9.0877 | -104.1403
9.0913 | -104.153
9.095 | -104.1681
9.0987 | -104.192
9.1023 | -104.2161
9.106 | -104.2441
9.1097 | -104.2716
9.1134 | -104.2929
9.117 | -104.3089
9.1207 | -104.3269
9.1244 | -104.34
9.1281 | -104.3489
9.1317 | -104.3492
9.1354 | -104.343
9.1391 | -104.3368
9.1428 | -104.3231
9.1464 | -104.3085
9.1501 | -104.2894
9.1538 | -104.2686
9.1574 | -104.2432
9.1611 | -104.216
9.1648 | -104.1931
9.1685 | -104.1753
9.1721 | -104.1537
9.1758 | -104.1367
9.1795 | -104.1219
9.1832 | -104.1127
9.1868 | -104.1125
9.1905 | -104.113
9.1942 | -104.1135
9.1979 | -104.1143
9.2015 | -104.1197
9.2052 | -104.1268
9.2089 | -104.1414
9.2125 | -104.1591
9.2162 | -104.1781
9.2199 | -104.1952
9.2236 | -104.2141
9.2272 | -104.237
9.2309 | -104.2634
9.2346 | -104.2857
9.2383 | -104.3077
9.2419 | -104.3309
9.2456 | -104.3505
9.2493 | -104.3735
9.253 | -104.3967
9.2566 | -104.4137
9.2603 | -104.4304
9.264 | -104.4445
9.2676 | -104.4574
9.2713 | -104.4669
9.275 | -104.4774
9.2787 | -104.486
9.2823 | -104.4873
9.286 | -104.4906
9.2897 | -104.4911
9.2934 | -104.4893
9.297 | -104.4896
9.3007 | -104.4912
9.3044 | -104.4911
9.308 | -104.4924
9.3117 | -104.4936
9.3154 | -104.4978
9.3191 | -104.4971
9.3227 | -104.496
9.3264 | -104.5006
9.3301 | -104.5052
9.3338 | -104.508
9.3374 | -104.5091
9.3411 | -104.5166
9.3448 | -104.5196
9.3485 | -104.5258
9.3521 | -104.5393
9.3558 | -104.5573
9.3595 | -104.5726
9.3631 | -104.5889
9.3668 | -104.6105
9.3705 | -104.6355
9.3742 | -104.66
9.3778 | -104.6824
9.3815 | -104.7049
9.3852 | -104.7211
9.3889 | -104.7309
9.3925 | -104.7437
9.3962 | -104.7496
9.3999 | -104.7511
9.4036 | -104.752
9.4072 | -104.7517
9.4109 | -104.7427
9.4146 | -104.7318
9.4182 | -104.7242
9.4219 | -104.7148
9.4256 | -104.7042
9.4293 | -104.6935
9.4329 | -104.6828
9.4366 | -104.6715
9.4403 | -104.6562
9.444 | -104.6422
9.4476 | -104.6319
9.4513 | -104.6235
9.455 | -104.6176
9.4587 | -104.6133
9.4623 | -104.6089
9.466 | -104.6063
9.4697 | -104.6065
9.4733 | -104.6043
9.477 | -104.6023
9.4807 | -104.6052
9.4844 | -104.611
9.488 | -104.617
9.4917 | -104.6249
9.4954 | -104.6349
9.4991 | -104.6423
9.5027 | -104.6497
9.5064 | -104.6565
9.5101 | -104.6641
9.5138 | -104.6707
9.5174 | -104.6742
9.5211 | -104.6795
9.5248 | -104.6831
9.5284 | -104.6753
9.5321 | -104.6728
9.5358 | -104.664
9.5395 | -104.6635
9.5431 | -104.6598
9.5468 | -104.6518
9.5505 | -104.6448
9.5542 | -104.6395
9.5578 | -104.6303
9.5615 | -104.6184
9.5652 | -104.6095
9.5689 | -104.6019
9.5725 | -104.5883
9.5762 | -104.5758
9.5799 | -104.5713
9.5835 | -104.5666
9.5872 | -104.5616
9.5909 | -104.558
9.5946 | -104.5549
9.5982 | -104.5538
9.6019 | -104.5566
9.6056 | -104.5632
9.6093 | -104.5719
9.6129 | -104.5876
9.6166 | -104.6044
9.6203 | -104.6217
9.624 | -104.6383
9.6276 | -104.6577
9.6313 | -104.6764
9.635 | -104.6962
9.6386 | -104.7188
9.6423 | -104.7399
9.646 | -104.7617
9.6497 | -104.7814
9.6533 | -104.7993
9.657 | -104.8143
9.6607 | -104.821
9.6644 | -104.8236
9.668 | -104.8247
9.6717 | -104.8179
9.6754 | -104.8102
9.679 | -104.7964
9.6827 | -104.7778
9.6864 | -104.7576
9.6901 | -104.7395
9.6937 | -104.7175
9.6974 | -104.6927
9.7011 | -104.6706
9.7048 | -104.6531
9.7084 | -104.6345
9.7121 | -104.6171
9.7158 | -104.604
9.7195 | -104.5926
9.7231 | -104.5835
9.7268 | -104.5792
9.7305 | -104.575
9.7341 | -104.5736
9.7378 | -104.5767
9.7415 | -104.58
9.7452 | -104.586
9.7488 | -104.5968
9.7525 | -104.6096
9.7562 | -104.6164
9.7599 | -104.6231
9.7635 | -104.6345
9.7672 | -104.6474
9.7709 | -104.6635
9.7746 | -104.684
9.7782 | -104.7024
9.7819 | -104.72
9.7856 | -104.7358
9.7892 | -104.7558
9.7929 | -104.7773
9.7966 | -104.7947
9.8003 | -104.8137
9.8039 | -104.8361
9.8076 | -104.8542
9.8113 | -104.8723
9.815 | -104.8901
9.8186 | -104.9026
9.8223 | -104.919
9.826 | -104.9324
9.8297 | -104.9438
9.8333 | -104.9558
9.837 | -104.968
9.8407 | -104.9805
9.8443 | -104.9879
9.848 | -104.9939
9.8517 | -105.001
9.8554 | -105.0075
9.859 | -105.0187
9.8627 | -105.0268
9.8664 | -105.0313
9.8701 | -105.0379
9.8737 | -105.0474
9.8774 | -105.0558
9.8811 | -105.0664
9.8848 | -105.0789
9.8884 | -105.0875
9.8921 | -105.0957
9.8958 | -105.1096
9.8994 | -105.1313
9.9031 | -105.1524
9.9068 | -105.167
9.9105 | -105.1819
9.9141 | -105.1954
9.9178 | -105.2132
9.9215 | -105.2319
9.9252 | -105.2452
9.9288 | -105.2583
9.9325 | -105.264
9.9362 | -105.2687
9.9398 | -105.2748
9.9435 | -105.2744
9.9472 | -105.2688
9.9509 | -105.2674
9.9545 | -105.2635
9.9582 | -105.2495
9.9619 | -105.2323
9.9656 | -105.2134
9.9692 | -105.1933
9.9729 | -105.1737
9.9766 | -105.1487
9.9803 | -105.1246
9.9839 | -105.1008
9.9876 | -105.078
9.9913 | -105.0526
9.9949 | -105.0307
9.9986 | -105.0128
10.0023 | -104.9923
10.006 | -104.9751
10.0096 | -104.9589
10.0133 | -104.9446
10.017 | -104.9349
10.0207 | -104.9305
10.0243 | -104.9248
10.028 | -104.9223
10.0317 | -104.9249
10.0354 | -104.9281
10.039 | -104.9265
10.0427 | -104.9232
10.0464 | -104.9215
10.05 | -104.9229
10.0537 | -104.9258
10.0574 | -104.9268
10.0611 | -104.923
10.0647 | -104.9185
10.0684 | -104.9159
10.0721 | -104.921
10.0758 | -104.9224
10.0794 | -104.9206
10.0831 | -104.9236
10.0868 | -104.9224
10.0905 | -104.9239
10.0941 | -104.9254
10.0978 | -104.9239
10.1015 | -104.9223
10.1051 | -104.9248
10.1088 | -104.9292
10.1125 | -104.9288
10.1162 | -104.927
10.1198 | -104.9241
10.1235 | -104.9242
10.1272 | -104.9242
10.1309 | -104.9235
10.1345 | -104.9258
10.1382 | -104.9325
10.1419 | -104.9399
10.1456 | -104.9447
10.1492 | -104.9517
10.1529 | -104.9616
10.1566 | -104.9742
10.1602 | -104.9896
10.1639 | -105.0041
10.1676 | -105.0151
10.1713 | -105.0306
10.1749 | -105.0446
10.1786 | -105.0538
10.1823 | -105.0638
10.186 | -105.0755
10.1896 | -105.0886
10.1933 | -105.0991
10.197 | -105.1091
10.2007 | -105.112
10.2043 | -105.1133
10.208 | -105.1176
10.2117 | -105.1239
10.2153 | -105.1239
10.219 | -105.1214
10.2227 | -105.117
10.2264 | -105.1155
10.23 | -105.1172
10.2337 | -105.1106
10.2374 | -105.1058
10.2411 | -105.0988
10.2447 | -105.0881
10.2484 | -105.0785
10.2521 | -105.0672
10.2558 | -105.057
10.2594 | -105.0478
10.2631 | -105.0392
10.2668 | -105.0336
10.2704 | -105.0283
10.2741 | -105.0221
10.2778 | -105.0164
10.2815 | -105.0081
10.2851 | -105.0007
10.2888 | -104.9931
10.2925 | -104.9871
10.2962 | -104.9852
10.2998 | -104.9834
10.3035 | -104.9795
10.3072 | -104.9733
10.3109 | -104.9629
10.3145 | -104.955
10.3182 | -104.951
10.3219 | -104.9463
10.3255 | -104.9426
10.3292 | -104.9416
10.3329 | -104.9392
10.3366 | -104.9401
10.3402 | -104.94
10.3439 | -104.9447
10.3476 | -104.9544
10.3513 | -104.9678
10.3549 | -104.9795
10.3586 | -104.9899
10.3623 | -105.003
10.3659 | -105.0242
10.3696 | -105.0424
10.3733 | -105.059
10.377 | -105.0759
10.3806 | -105.092
10.3843 | -105.1094
10.388 | -105.1243
10.3917 | -105.1391
10.3953 | -105.1536
10.399 | -105.1643
10.4027 | -105.1776
10.4064 | -105.1892
10.41 | -105.1975
10.4137 | -105.2037
10.4174 | -105.206
10.421 | -105.2093
10.4247 | -105.212
10.4284 | -105.2125
10.4321 | -105.2167
10.4357 | -105.217
10.4394 | -105.2161
10.4431 | -105.2116
10.4468 | -105.2054
10.4504 | -105.1959
10.4541 | -105.1885
10.4578 | -105.1793
10.4615 | -105.1667
10.4651 | -105.156
10.4688 | -105.1434
10.4725 | -105.1299
10.4761 | -105.1198
10.4798 | -105.1098
10.4835 | -105.0981
10.4872 | -105.0824
10.4908 | -105.0697
10.4945 | -105.0577
10.4982 | -105.0458
10.5019 | -105.0407
10.5055 | -105.0387
10.5092 | -105.0386
10.5129 | -105.0389
10.5166 | -105.0386
10.5202 | -105.0429
10.5239 | -105.0546
10.5276 | -105.0694
10.5312 | -105.0845
10.5349 | -105.101
10.5386 | -105.1226
10.5423 | -105.1431
10.5459 | -105.1636
10.5496 | -105.1818
10.5533 | -105.1972
10.557 | -105.2108
10.5606 | -105.2226
10.5643 | -105.2287
10.568 | -105.2322
10.5717 | -105.2333
10.5753 | -105.2314
10.579 | -105.2249
10.5827 | -105.2142
10.5863 | -105.1991
10.59 | -105.1783
10.5937 | -105.1583
10.5974 | -105.1321
10.601 | -105.1107
10.6047 | -105.0889
10.6084 | -105.0646
10.6121 | -105.0419
10.6157 | -105.0188
10.6194 | -104.9944
10.6231 | -104.9741
10.6268 | -104.9558
10.6304 | -104.9395
10.6341 | -104.9284
10.6378 | -104.9229
10.6414 | -104.9192
10.6451 | -104.9172
10.6488 | -104.92
10.6525 | -104.9239
10.6561 | -104.928
10.6598 | -104.933
10.6635 | -104.9408
10.6672 | -104.9506
10.6708 | -104.96
10.6745 | -104.9664
10.6782 | -104.9707
10.6818 | -104.9763
10.6855 | -104.9834
10.6892 | -104.9894
10.6929 | -104.9924
10.6965 | -104.9945
10.7002 | -104.9982
10.7039 | -105.0034
10.7076 | -105.0079
10.7112 | -105.0101
10.7149 | -105.0134
10.7186 | -105.0192
10.7223 | -105.023
10.7259 | -105.0225
10.7296 | -105.0236
10.7333 | -105.029
10.7369 | -105.0344
10.7406 | -105.0473
10.7443 | -105.0621
10.748 | -105.073
10.7516 | -105.086
10.7553 | -105.1016
10.759 | -105.1198
10.7627 | -105.1424
10.7663 | -105.1663
10.77 | -105.1917
10.7737 | -105.2178
10.7774 | -105.2372
10.781 | -105.2588
10.7847 | -105.2793
10.7884 | -105.2939
10.792 | -105.305
10.7957 | -105.3128
10.7994 | -105.3193
10.8031 | -105.3219
10.8067 | -105.3173
10.8104 | -105.3095
10.8141 | -105.3017
10.8178 | -105.2922
10.8214 | -105.2789
10.8251 | -105.2648
10.8288 | -105.2507
10.8325 | -105.2392
10.8361 | -105.2279
10.8398 | -105.2149
10.8435 | -105.2007
10.8471 | -105.1918
10.8508 | -105.1855
10.8545 | -105.1827
10.8582 | -105.179
10.8618 | -105.1756
10.8655 | -105.172
10.8692 | -105.1715
10.8729 | -105.1719
10.8765 | -105.1712
10.8802 | -105.1709
10.8839 | -105.1716
10.8876 | -105.1725
10.8912 | -105.1725
10.8949 | -105.1692
10.8986 | -105.1596
10.9022 | -105.1536
10.9059 | -105.152
10.9096 | -105.1475
10.9133 | -105.1433
10.9169 | -105.1382
10.9206 | -105.1324
10.9243 | -105.126
10.928 | -105.1214
10.9316 | -105.116
10.9353 | -105.1099
10.939 | -105.0995
10.9427 | -105.0921
10.9463 | -105.089
10.95 | -105.086
10.9537 | -105.0849
10.9573 | -105.0843
10.961 | -105.0867
10.9647 | -105.0911
10.9684 | -105.0969
10.972 | -105.1033
10.9757 | -105.111
10.9794 | -105.1233
10.9831 | -105.1344
10.9867 | -105.1474
10.9904 | -105.1629
10.9941 | -105.1781
10.9977 | -105.1955
11.0014 | -105.2182
11.0051 | -105.2363
11.0088 | -105.254
11.0124 | -105.2745
11.0161 | -105.2979
11.0198 | -105.3241
11.0235 | -105.3506
11.0271 | -105.3781
11.0308 | -105.4065
11.0345 | -105.4384
11.0382 | -105.4703
11.0418 | -105.5006
11.0455 | -105.5307
11.0492 | -105.5588
11.0528 | -105.5897
11.0565 | -105.6252
11.0602 | -105.6571
11.0639 | -105.6827
11.0675 | -105.7036
11.0712 | -105.7238
11.0749 | -105.7393
11.0786 | -105.747
11.0822 | -105.7556
11.0859 | -105.7636
11.0896 | -105.7708
11.0933 | -105.7752
11.0969 | -105.7743
11.1006 | -105.7689
11.1043 | -105.7584
11.1079 | -105.7456
11.1116 | -105.7308
11.1153 | -105.7149
11.119 | -105.6969
11.1226 | -105.6809
11.1263 | -105.6666
11.13 | -105.649
11.1337 | -105.6335
11.1373 | -105.6219
11.141 | -105.6053
11.1447 | -105.5948
11.1484 | -105.5876
11.152 | -105.5791
11.1557 | -105.5683
11.1594 | -105.5601
11.163 | -105.5527
11.1667 | -105.551
11.1704 | -105.5519
11.1741 | -105.5486
11.1777 | -105.5469
11.1814 | -105.5513
11.1851 | -105.5577
11.1888 | -105.5638
11.1924 | -105.5722
11.1961 | -105.5788
11.1998 | -105.59
11.2035 | -105.6023
11.2071 | -105.6119
11.2108 | -105.6237
11.2145 | -105.636
11.2181 | -105.6418
11.2218 | -105.6477
11.2255 | -105.6539
11.2292 | -105.66
11.2328 | -105.6598
11.2365 | -105.6581
11.2402 | -105.6539
11.2439 | -105.6418
11.2475 | -105.6243
11.2512 | -105.6058
11.2549 | -105.5869
11.2586 | -105.5661
11.2622 | -105.5437
11.2659 | -105.5196
11.2696 | -105.492
11.2732 | -105.462
11.2769 | -105.4342
11.2806 | -105.4072
11.2843 | -105.378
11.2879 | -105.347
11.2916 | -105.3223
11.2953 | -105.3024
11.299 | -105.2836
11.3026 | -105.2633
11.3063 | -105.2461
11.31 | -105.2306
11.3137 | -105.2198
11.3173 | -105.2121
11.321 | -105.2105
11.3247 | -105.2095
11.3283 | -105.2068
11.332 | -105.2067
11.3357 | -105.2092
11.3394 | -105.2171
11.343 | -105.2234
11.3467 | -105.2263
11.3504 | -105.232
11.3541 | -105.241
11.3577 | -105.2461
11.3614 | -105.2506
11.3651 | -105.2534
11.3688 | -105.2605
11.3724 | -105.2705
11.3761 | -105.2773
11.3798 | -105.2813
11.3834 | -105.2883
11.3871 | -105.2941
11.3908 | -105.2993
11.3945 | -105.3082
11.3981 | -105.3175
11.4018 | -105.3249
11.4055 | -105.3354
11.4092 | -105.3448
11.4128 | -105.352
11.4165 | -105.3601
11.4202 | -105.3672
11.4238 | -105.3711
11.4275 | -105.3744
11.4312 | -105.3767
11.4349 | -105.3789
11.4385 | -105.3804
11.4422 | -105.3792
11.4459 | -105.3795
11.4496 | -105.3787
11.4532 | -105.3746
11.4569 | -105.3686
11.4606 | -105.3617
11.4643 | -105.3533
11.4679 | -105.3463
11.4716 | -105.3394
11.4753 | -105.3311
11.4789 | -105.3213
11.4826 | -105.311
11.4863 | -105.3016
11.49 | -105.2912
11.4936 | -105.2841
11.4973 | -105.2748
11.501 | -105.2657
11.5047 | -105.2542
11.5083 | -105.2482
11.512 | -105.2445
11.5157 | -105.2411
11.5194 | -105.2385
11.523 | -105.2421
11.5267 | -105.249
11.5304 | -105.2572
11.534 | -105.2655
11.5377 | -105.275
11.5414 | -105.2916
11.5451 | -105.311
11.5487 | -105.3317
11.5524 | -105.3536
11.5561 | -105.3816
11.5598 | -105.407
11.5634 | -105.4344
11.5671 | -105.4668
11.5708 | -105.4985
11.5745 | -105.5345
11.5781 | -105.5742
11.5818 | -105.6098
11.5855 | -105.6474
11.5891 | -105.6866
11.5928 | -105.7211
11.5965 | -105.7551
11.6002 | -105.7892
11.6038 | -105.8201
11.6075 | -105.8415
11.6112 | -105.8561
11.6149 | -105.8732
11.6185 | -105.8889
11.6222 | -105.8989
11.6259 | -105.9014
11.6296 | -105.8953
11.6332 | -105.8857
11.6369 | -105.8743
11.6406 | -105.8566
11.6442 | -105.834
11.6479 | -105.8139
11.6516 | -105.7841
11.6553 | -105.7437
11.6589 | -105.715
11.6626 | -105.68
11.6663 | -105.6455
11.67 | -105.6095
11.6736 | -105.5781
11.6773 | -105.5435
11.681 | -105.5056
11.6846 | -105.4729
11.6883 | -105.4413
11.692 | -105.4097
11.6957 | -105.3821
11.6993 | -105.3569
11.703 | -105.3342
11.7067 | -105.3152
11.7104 | -105.2968
11.714 | -105.2791
11.7177 | -105.2689
11.7214 | -105.2644
11.7251 | -105.2589
11.7287 | -105.2582
11.7324 | -105.2635
11.7361 | -105.2726
11.7397 | -105.2824
11.7434 | -105.2928
11.7471 | -105.3098
11.7508 | -105.332
11.7544 | -105.3562
11.7581 | -105.382
11.7618 | -105.4115
11.7655 | -105.4416
11.7691 | -105.4713
11.7728 | -105.5023
11.7765 | -105.5348
11.7802 | -105.5683
11.7838 | -105.5976
11.7875 | -105.6233
11.7912 | -105.6434
11.7948 | -105.6653
11.7985 | -105.6907
11.8022 | -105.7123
11.8059 | -105.7289
11.8095 | -105.7408
11.8132 | -105.7487
11.8169 | -105.7501
11.8206 | -105.7579
11.8242 | -105.7616
11.8279 | -105.7549
11.8316 | -105.7444
11.8353 | -105.7362
11.8389 | -105.7219
11.8426 | -105.7032
11.8463 | -105.682
11.8499 | -105.6575
11.8536 | -105.6325
11.8573 | -105.5992
11.861 | -105.565
11.8646 | -105.5312
11.8683 | -105.4958
11.872 | -105.46
11.8757 | -105.4223
11.8793 | -105.3855
11.883 | -105.353
11.8867 | -105.324
11.8904 | -105.2915
11.894 | -105.2609
11.8977 | -105.2349
11.9014 | -105.2105
11.905 | -105.1903
11.9087 | -105.1717
11.9124 | -105.1563
11.9161 | -105.1469
11.9197 | -105.1345
11.9234 | -105.1277
11.9271 | -105.1271
11.9308 | -105.1294
11.9344 | -105.1319
11.9381 | -105.1385
11.9418 | -105.1485
11.9455 | -105.1596
11.9491 | -105.1761
11.9528 | -105.1934
11.9565 | -105.2127
11.9601 | -105.2369
11.9638 | -105.2612
11.9675 | -105.2852
11.9712 | -105.3086
11.9748 | -105.3309
11.9785 | -105.3568
11.9822 | -105.3858
11.9859 | -105.4138
11.9895 | -105.4411
11.9932 | -105.4638
11.9969 | -105.4801
12.0006 | -105.4955
12.0042 | -105.5176
12.0079 | -105.5339
12.0116 | -105.5487
12.0152 | -105.5624
12.0189 | -105.5694
12.0226 | -105.5796
12.0263 | -105.5909
12.0299 | -105.6031
12.0336 | -105.615
12.0373 | -105.6253
12.041 | -105.6338
12.0446 | -105.6457
12.0483 | -105.6599
12.052 | -105.6741
12.0556 | -105.6878
12.0593 | -105.7022
12.063 | -105.7153
12.0667 | -105.731
12.0703 | -105.7494
12.074 | -105.7675
12.0777 | -105.7813
12.0814 | -105.7904
12.085 | -105.798
12.0887 | -105.8041
12.0924 | -105.8126
12.0961 | -105.8151
12.0997 | -105.8126
12.1034 | -105.8045
12.1071 | -105.7956
12.1107 | -105.7842
12.1144 | -105.7669
12.1181 | -105.7424
12.1218 | -105.7055
12.1254 | -105.6699
12.1291 | -105.6345
12.1328 | -105.6
12.1365 | -105.5674
12.1401 | -105.5293
12.1438 | -105.4926
12.1475 | -105.4613
12.1512 | -105.4271
12.1548 | -105.3947
12.1585 | -105.3661
12.1622 | -105.3405
12.1658 | -105.3221
12.1695 | -105.3138
12.1732 | -105.3073
12.1769 | -105.2976
12.1805 | -105.2958
12.1842 | -105.2992
12.1879 | -105.3251
12.1916 | -105.3355
12.1952 | -105.3504
12.1989 | -105.3652
12.2026 | -105.3821
12.2063 | -105.4031
12.2099 | -105.4247
12.2136 | -105.4479
12.2173 | -105.475
12.2209 | -105.4982
12.2246 | -105.5223
12.2283 | -105.5486
12.232 | -105.576
12.2356 | -105.6011
12.2393 | -105.6267
12.243 | -105.6486
12.2467 | -105.6686
12.2503 | -105.6911
12.254 | -105.7094
12.2577 | -105.7295
12.2614 | -105.7499
12.265 | -105.7634
12.2687 | -105.7742
12.2724 | -105.7844
12.276 | -105.7921
12.2797 | -105.8013
12.2834 | -105.8102
12.2871 | -105.8165
12.2907 | -105.8252
12.2944 | -105.8335
12.2981 | -105.8457
12.3018 | -105.8529
12.3054 | -105.8563
12.3091 | -105.8576
12.3128 | -105.8613
12.3165 | -105.8672
12.3201 | -105.8717
12.3238 | -105.8753
12.3275 | -105.8796
12.3311 | -105.8842
12.3348 | -105.8929
12.3385 | -105.9017
12.3422 | -105.9065
12.3458 | -105.9096
12.3495 | -105.9129
12.3532 | -105.9182
12.3569 | -105.9193
12.3605 | -105.9204
12.3642 | -105.9162
12.3679 | -105.9154
12.3716 | -105.9141
12.3752 | -105.9097
12.3789 | -105.9059
12.3826 | -105.9003
12.3862 | -105.8922
12.3899 | -105.8832
12.3936 | -105.8728
12.3973 | -105.8625
12.4009 | -105.8557
12.4046 | -105.8458
12.4083 | -105.8339
12.412 | -105.8192
12.4156 | -105.8081
12.4193 | -105.7994
12.423 | -105.7925
12.4267 | -105.7841
12.4303 | -105.7727
12.434 | -105.7669
12.4377 | -105.7644
12.4413 | -105.7629
12.445 | -105.7662
12.4487 | -105.7714
12.4524 | -105.7797
12.456 | -105.7874
12.4597 | -105.796
12.4634 | -105.8059
12.4671 | -105.8192
12.4707 | -105.8327
12.4744 | -105.8501
12.4781 | -105.8689
12.4817 | -105.8872
12.4854 | -105.905
12.4891 | -105.9229
12.4928 | -105.9414
12.4964 | -105.9595
12.5001 | -105.9744
12.5038 | -105.986
12.5075 | -105.9996
12.5111 | -106.0114
12.5148 | -106.0196
12.5185 | -106.0238
12.5222 | -106.0293
12.5258 | -106.0343
12.5295 | -106.035
12.5332 | -106.0342
12.5368 | -106.0361
12.5405 | -106.0374
12.5442 | -106.0376
12.5479 | -106.0307
12.5515 | -106.0208
12.5552 | -106.0108
12.5589 | -105.9985
12.5626 | -105.985
12.5662 | -105.9697
12.5699 | -105.9567
12.5736 | -105.9425
12.5773 | -105.9253
12.5809 | -105.9062
12.5846 | -105.8872
12.5883 | -105.8668
12.5919 | -105.8485
12.5956 | -105.8297
12.5993 | -105.8079
12.603 | -105.7849
12.6066 | -105.7652
12.6103 | -105.7488
12.614 | -105.733
12.6177 | -105.7177
12.6213 | -105.7006
12.625 | -105.6848
12.6287 | -105.6703
12.6324 | -105.6602
12.636 | -105.6561
12.6397 | -105.6519
12.6434 | -105.6476
12.647 | -105.6452
12.6507 | -105.6447
12.6544 | -105.6468
12.6581 | -105.6535
12.6617 | -105.66
12.6654 | -105.6687
12.6691 | -105.6806
12.6728 | -105.6942
12.6764 | -105.709
12.6801 | -105.7281
12.6838 | -105.7499
12.6875 | -105.7716
12.6911 | -105.7954
12.6948 | -105.819
12.6985 | -105.845
12.7021 | -105.8674
12.7058 | -105.8903
12.7095 | -105.9161
12.7132 | -105.9403
12.7168 | -105.9599
12.7205 | -105.98
12.7242 | -105.9992
12.7279 | -106.0198
12.7315 | -106.0361
12.7352 | -106.0554
12.7389 | -106.0736
12.7425 | -106.0849
12.7462 | -106.0914
12.7499 | -106.0985
12.7536 | -106.1027
12.7572 | -106.1081
12.7609 | -106.1108
12.7646 | -106.1118
12.7683 | -106.1134
12.7719 | -106.1122
12.7756 | -106.1087
12.7793 | -106.1045
12.783 | -106.102
12.7866 | -106.0977
12.7903 | -106.0922
12.794 | -106.0846
12.7976 | -106.076
12.8013 | -106.0685
12.805 | -106.0635
12.8087 | -106.0561
12.8123 | -106.0503
12.816 | -106.0409
12.8197 | -106.0324
12.8234 | -106.0271
12.827 | -106.0226
12.8307 | -106.0171
12.8344 | -106.0147
12.8381 | -106.0128
12.8417 | -106.0087
12.8454 | -106.0069
12.8491 | -106.0092
12.8527 | -106.0097
12.8564 | -106.0118
12.8601 | -106.0169
12.8638 | -106.0206
12.8674 | -106.0253
12.8711 | -106.0271
12.8748 | -106.0281
12.8785 | -106.0296
12.8821 | -106.0301
12.8858 | -106.0329
12.8895 | -106.035
12.8932 | -106.035
12.8968 | -106.0351
12.9005 | -106.0344
12.9042 | -106.0311
12.9078 | -106.0277
12.9115 | -106.0263
12.9152 | -106.0243
12.9189 | -106.0239
12.9225 | -106.0215
12.9262 | -106.0197
12.9299 | -106.0204
12.9336 | -106.023
12.9372 | -106.0227
12.9409 | -106.0238
12.9446 | -106.0247
12.9483 | -106.0269
12.9519 | -106.0328
12.9556 | -106.0416
12.9593 | -106.0535
12.9629 | -106.0665
12.9666 | -106.0804
12.9703 | -106.0969
12.974 | -106.1164
12.9776 | -106.1366
12.9813 | -106.1614
12.985 | -106.1836
12.9887 | -106.2041
12.9923 | -106.2264
12.996 | -106.2517
12.9997 | -106.2735
13.0034 | -106.299
13.007 | -106.3246
13.0107 | -106.3493
13.0144 | -106.3696
13.018 | -106.3868
13.0217 | -106.3987
13.0254 | -106.4093
13.0291 | -106.4173
13.0327 | -106.4227
13.0364 | -106.4227
13.0401 | -106.4208
13.0438 | -106.4166
13.0474 | -106.4102
13.0511 | -106.3982
13.0548 | -106.3768
13.0585 | -106.3596
13.0621 | -106.3396
13.0658 | -106.3169
13.0695 | -106.2918
13.0731 | -106.2575
13.0768 | -106.22
13.0805 | -106.1828
13.0842 | -106.1449
13.0878 | -106.1088
13.0915 | -106.0727
13.0952 | -106.0363
13.0989 | -105.9967
13.1025 | -105.9574
13.1062 | -105.9239
13.1099 | -105.8927
13.1136 | -105.8634
13.1172 | -105.8365
13.1209 | -105.8111
13.1246 | -105.7931
13.1282 | -105.7776
13.1319 | -105.7639
13.1356 | -105.7547
13.1393 | -105.7491
13.1429 | -105.7434
13.1466 | -105.7385
13.1503 | -105.7375
13.154 | -105.7378
13.1576 | -105.7369
13.1613 | -105.7399
13.165 | -105.7443
13.1686 | -105.7481
13.1723 | -105.754
13.176 | -105.7618
13.1797 | -105.7702
13.1833 | -105.778
13.187 | -105.7851
13.1907 | -105.7899
13.1944 | -105.7969
13.198 | -105.8044
13.2017 | -105.8153
13.2054 | -105.8259
13.2091 | -105.8373
13.2127 | -105.8448
13.2164 | -105.8526
13.2201 | -105.861
13.2237 | -105.8694
13.2274 | -105.8763
13.2311 | -105.88
13.2348 | -105.8825
13.2384 | -105.8912
13.2421 | -105.9017
13.2458 | -105.9035
13.2495 | -105.906
13.2531 | -105.9081
13.2568 | -105.9123
13.2605 | -105.9192
13.2642 | -105.9285
13.2678 | -105.9386
13.2715 | -105.9472
13.2752 | -105.9559
13.2788 | -105.9675
13.2825 | -105.9787
13.2862 | -105.99
13.2899 | -106.0032
13.2935 | -106.0194
13.2972 | -106.0323
13.3009 | -106.0445
13.3046 | -106.0562
13.3082 | -106.0693
13.3119 | -106.0827
13.3156 | -106.0946
13.3193 | -106.1027
13.3229 | -106.1123
13.3266 | -106.1191
13.3303 | -106.1251
13.3339 | -106.1312
13.3376 | -106.1341
13.3413 | -106.1343
13.345 | -106.137
13.3486 | -106.14
13.3523 | -106.1383
13.356 | -106.1352
13.3597 | -106.1286
13.3633 | -106.1234
13.367 | -106.1167
13.3707 | -106.1087
13.3744 | -106.0987
13.378 | -106.091
13.3817 | -106.0844
13.3854 | -106.0716
13.389 | -106.0559
13.3927 | -106.0391
13.3964 | -106.0272
13.4001 | -106.0169
13.4037 | -106.0073
13.4074 | -105.9942
13.4111 | -105.9795
13.4148 | -105.9642
13.4184 | -105.952
13.4221 | -105.9421
13.4258 | -105.9332
13.4294 | -105.9221
13.4331 | -105.9138
13.4368 | -105.9095
13.4405 | -105.9056
13.4441 | -105.8995
13.4478 | -105.8957
13.4515 | -105.8904
13.4552 | -105.8881
13.4588 | -105.8892
13.4625 | -105.8897
13.4662 | -105.8921
13.4699 | -105.8955
13.4735 | -105.8991
13.4772 | -105.9091
13.4809 | -105.9216
13.4845 | -105.9361
13.4882 | -105.9486
13.4919 | -105.9616
13.4956 | -105.9771
13.4992 | -105.9938
13.5029 | -106.0108
13.5066 | -106.0297
13.5103 | -106.0484
13.5139 | -106.0681
13.5176 | -106.0844
13.5213 | -106.1045
13.525 | -106.121
13.5286 | -106.137
13.5323 | -106.1568
13.536 | -106.1746
13.5396 | -106.1882
13.5433 | -106.2002
13.547 | -106.2107
13.5507 | -106.2213
13.5543 | -106.2296
13.558 | -106.2357
13.5617 | -106.2402
13.5654 | -106.2434
13.569 | -106.2482
13.5727 | -106.2507
13.5764 | -106.2522
13.5801 | -106.2546
13.5837 | -106.2562
13.5874 | -106.2574
13.5911 | -106.2538
13.5947 | -106.2475
13.5984 | -106.2423
13.6021 | -106.2377
13.6058 | -106.2328
13.6094 | -106.2281
13.6131 | -106.2202
13.6168 | -106.209
13.6205 | -106.1987
13.6241 | -106.1919
13.6278 | -106.1853
13.6315 | -106.1777
13.6352 | -106.1696
13.6388 | -106.1658
13.6425 | -106.1616
13.6462 | -106.1582
13.6498 | -106.1554
13.6535 | -106.1546
13.6572 | -106.1557
13.6609 | -106.1576
13.6645 | -106.1561
13.6682 | -106.1558
13.6719 | -106.1587
13.6756 | -106.1653
13.6792 | -106.1704
13.6829 | -106.1777
13.6866 | -106.1845
13.6903 | -106.1889
13.6939 | -106.1923
13.6976 | -106.1923
13.7013 | -106.19
13.7049 | -106.1866
13.7086 | -106.1855
13.7123 | -106.1846
13.716 | -106.179
13.7196 | -106.1711
13.7233 | -106.1601
13.727 | -106.1465
13.7307 | -106.1296
13.7343 | -106.1129
13.738 | -106.0951
13.7417 | -106.0759
13.7454 | -106.059
13.749 | -106.0467
13.7527 | -106.0303
13.7564 | -106.0132
13.76 | -105.9946
13.7637 | -105.9761
13.7674 | -105.9587
13.7711 | -105.9418
13.7747 | -105.9353
13.7784 | -105.9249
13.7821 | -105.9167
13.7858 | -105.9083
13.7894 | -105.904
13.7931 | -105.9018
13.7968 | -105.8999
13.8004 | -105.8995
13.8041 | -105.9015
13.8078 | -105.905
13.8115 | -105.9068
13.8151 | -105.9094
13.8188 | -105.9125
13.8225 | -105.9164
13.8262 | -105.9206
13.8298 | -105.9271
13.8335 | -105.9361
13.8372 | -105.9458
13.8409 | -105.9561
13.8445 | -105.966
13.8482 | -105.9726
13.8519 | -105.9806
13.8555 | -105.9925
13.8592 | -106.0044
13.8629 | -106.0194
13.8666 | -106.0331
13.8702 | -106.0433
13.8739 | -106.0516
13.8776 | -106.0612
13.8813 | -106.0731
13.8849 | -106.0843
13.8886 | -106.0948
13.8923 | -106.1064
13.896 | -106.1205
13.8996 | -106.1317
13.9033 | -106.1416
13.907 | -106.1507
13.9106 | -106.1584
13.9143 | -106.1649
13.918 | -106.1722
13.9217 | -106.1776
13.9253 | -106.1767
13.929 | -106.1767
13.9327 | -106.1784
13.9364 | -106.1806
13.94 | -106.181
13.9437 | -106.1748
13.9474 | -106.1644
13.9511 | -106.1522
13.9547 | -106.142
13.9584 | -106.1311
13.9621 | -106.1152
13.9657 | -106.0982
13.9694 | -106.0801
13.9731 | -106.0567
13.9768 | -106.0323
13.9804 | -106.0076
13.9841 | -105.9821
13.9878 | -105.9543
13.9915 | -105.9246
13.9951 | -105.8962
13.9988 | -105.869
14.0025 | -105.8423
14.0062 | -105.8148
14.0098 | -105.786
14.0135 | -105.7601
14.0172 | -105.737
14.0208 | -105.7148
14.0245 | -105.6926
14.0282 | -105.6734
14.0319 | -105.6543
14.0355 | -105.6387
14.0392 | -105.627
14.0429 | -105.6185
14.0466 | -105.618
14.0502 | -105.6226
14.0539 | -105.6265
14.0576 | -105.632
14.0613 | -105.6448
14.0649 | -105.6616
14.0686 | -105.6811
14.0723 | -105.702
14.0759 | -105.7235
14.0796 | -105.7477
14.0833 | -105.773
14.087 | -105.8009
14.0906 | -105.8316
14.0943 | -105.8597
14.098 | -105.8883
14.1017 | -105.918
14.1053 | -105.9459
14.109 | -105.9708
14.1127 | -105.993
14.1164 | -106.0141
14.12 | -106.0326
14.1237 | -106.0503
14.1274 | -106.0691
14.131 | -106.0859
14.1347 | -106.0958
14.1384 | -106.1033
14.1421 | -106.1111
14.1457 | -106.1182
14.1494 | -106.1252
14.1531 | -106.1336
14.1568 | -106.1393
14.1604 | -106.1414
14.1641 | -106.1465
14.1678 | -106.1537
14.1715 | -106.1594
14.1751 | -106.1687
14.1788 | -106.1751
14.1825 | -106.1807
14.1861 | -106.1892
14.1898 | -106.1986
14.1935 | -106.2042
14.1972 | -106.2115
14.2008 | -106.2196
14.2045 | -106.2293
14.2082 | -106.242
14.2119 | -106.2553
14.2155 | -106.2709
14.2192 | -106.2846
14.2229 | -106.297
14.2265 | -106.3097
14.2302 | -106.3206
14.2339 | -106.3353
14.2376 | -106.3456
14.2412 | -106.3549
14.2449 | -106.3642
14.2486 | -106.3687
14.2523 | -106.3727
14.2559 | -106.3808
14.2596 | -106.3842
14.2633 | -106.3829
14.267 | -106.3823
14.2706 | -106.3801
14.2743 | -106.3785
14.278 | -106.376
14.2816 | -106.3701
14.2853 | -106.3623
14.289 | -106.3528
14.2927 | -106.3399
14.2963 | -106.3282
14.3 | -106.3148
14.3037 | -106.3
14.3074 | -106.285
14.311 | -106.2716
14.3147 | -106.2573
14.3184 | -106.2432
14.3221 | -106.2306
14.3257 | -106.2183
14.3294 | -106.2077
14.3331 | -106.1932
14.3367 | -106.1779
14.3404 | -106.1664
14.3441 | -106.155
14.3478 | -106.1454
14.3514 | -106.1386
14.3551 | -106.1321
14.3588 | -106.1255
14.3625 | -106.1171
14.3661 | -106.1128
14.3698 | -106.1105
14.3735 | -106.1058
14.3772 | -106.1012
14.3808 | -106.0996
14.3845 | -106.099
14.3882 | -106.0989
14.3918 | -106.0975
14.3955 | -106.1003
14.3992 | -106.1092
14.4029 | -106.1158
14.4065 | -106.1198
14.4102 | -106.1231
14.4139 | -106.1244
14.4176 | -106.127
14.4212 | -106.1263
14.4249 | -106.1232
14.4286 | -106.1218
14.4323 | -106.1211
14.4359 | -106.1202
14.4396 | -106.1196
14.4433 | -106.1164
14.4469 | -106.1101
14.4506 | -106.1052
14.4543 | -106.1019
14.458 | -106.0975
14.4616 | -106.0902
14.4653 | -106.0831
14.469 | -106.0758
14.4727 | -106.0653
14.4763 | -106.0568
14.48 | -106.0501
14.4837 | -106.0437
14.4873 | -106.0363
14.491 | -106.0303
14.4947 | -106.027
14.4984 | -106.0263
14.502 | -106.0263
14.5057 | -106.0294
14.5094 | -106.0333
14.5131 | -106.0395
14.5167 | -106.0493
14.5204 | -106.0586
14.5241 | -106.0674
14.5278 | -106.0797
14.5314 | -106.0909
14.5351 | -106.1035
14.5388 | -106.1166
14.5424 | -106.1329
14.5461 | -106.148
14.5498 | -106.1615
14.5535 | -106.1755
14.5571 | -106.1897
14.5608 | -106.2026
14.5645 | -106.2149
14.5682 | -106.2271
14.5718 | -106.239
14.5755 | -106.2487
14.5792 | -106.2524
14.5829 | -106.2569
14.5865 | -106.2619
14.5902 | -106.2623
14.5939 | -106.2625
14.5975 | -106.2606
14.6012 | -106.2613
14.6049 | -106.2589
14.6086 | -106.2514
14.6122 | -106.2441
14.6159 | -106.2388
14.6196 | -106.2347
14.6233 | -106.2315
14.6269 | -106.2298
14.6306 | -106.2299
14.6343 | -106.2311
14.638 | -106.2284
14.6416 | -106.2246
14.6453 | -106.2252
14.649 | -106.2208
14.6526 | -106.2213
14.6563 | -106.2257
14.66 | -106.2288
14.6637 | -106.2299
14.6673 | -106.2326
14.671 | -106.2384
14.6747 | -106.2469
14.6784 | -106.2572
14.682 | -106.2695
14.6857 | -106.2815
14.6894 | -106.2902
14.6931 | -106.3008
14.6967 | -106.3148
14.7004 | -106.3266
14.7041 | -106.3365
14.7077 | -106.3485
14.7114 | -106.3617
14.7151 | -106.3726
14.7188 | -106.3837
14.7224 | -106.3942
14.7261 | -106.4027
14.7298 | -106.414
14.7335 | -106.4291
14.7371 | -106.4427
14.7408 | -106.4538
14.7445 | -106.4636
14.7482 | -106.4728
14.7518 | -106.484
14.7555 | -106.496
14.7592 | -106.5073
14.7628 | -106.5173
14.7665 | -106.5256
14.7702 | -106.532
14.7739 | -106.5393
14.7775 | -106.5464
14.7812 | -106.5529
14.7849 | -106.5592
14.7886 | -106.5679
14.7922 | -106.573
14.7959 | -106.576
14.7996 | -106.5784
14.8033 | -106.5762
14.8069 | -106.5747
14.8106 | -106.5695
14.8143 | -106.5597
14.8179 | -106.5526
14.8216 | -106.5432
14.8253 | -106.5333
14.829 | -106.5211
14.8326 | -106.5049
14.8363 | -106.4883
14.84 | -106.4698
14.8437 | -106.4542
14.8473 | -106.4408
14.851 | -106.428
14.8547 | -106.4148
14.8583 | -106.3967
14.862 | -106.384
14.8657 | -106.371
14.8694 | -106.3574
14.873 | -106.3414
14.8767 | -106.3265
14.8804 | -106.3119
14.8841 | -106.2974
14.8877 | -106.283
14.8914 | -106.269
14.8951 | -106.256
14.8988 | -106.2446
14.9024 | -106.2355
14.9061 | -106.2261
14.9098 | -106.215
14.9134 | -106.2053
14.9171 | -106.1979
14.9208 | -106.1911
14.9245 | -106.1864
14.9281 | -106.1836
14.9318 | -106.1809
14.9355 | -106.1784
14.9392 | -106.178
14.9428 | -106.1784
14.9465 | -106.1785
14.9502 | -106.178
14.9539 | -106.1785
14.9575 | -106.1807
14.9612 | -106.1827
14.9649 | -106.1825
14.9685 | -106.1808
14.9722 | -106.1803
14.9759 | -106.1813
14.9796 | -106.1811
14.9832 | -106.1799
14.9869 | -106.174
14.9906 | -106.1698
14.9943 | -106.1662
14.9979 | -106.1623
15.0016 | -106.1534
15.0053 | -106.1458
15.009 | -106.1387
15.0126 | -106.1314
15.0163 | -106.1245
15.02 | -106.1155
15.0236 | -106.1055
15.0273 | -106.0958
15.031 | -106.0843
15.0347 | -106.074
15.0383 | -106.0634
15.042 | -106.0523
15.0457 | -106.0437
15.0494 | -106.0346
15.053 | -106.0247
15.0567 | -106.0193
15.0604 | -106.014
15.0641 | -106.0055
15.0677 | -105.9992
15.0714 | -105.9963
15.0751 | -105.9982
15.0787 | -106.0002
15.0824 | -106.0053
15.0861 | -106.0121
15.0898 | -106.0238
15.0934 | -106.0383
15.0971 | -106.0551
15.1008 | -106.0733
15.1045 | -106.0922
15.1081 | -106.1121
15.1118 | -106.1337
15.1155 | -106.1572
15.1192 | -106.1815
15.1228 | -106.2067
15.1265 | -106.2356
15.1302 | -106.2637
15.1338 | -106.2936
15.1375 | -106.3246
15.1412 | -106.3552
15.1449 | -106.386
15.1485 | -106.4169
15.1522 | -106.4464
15.1559 | -106.4719
15.1596 | -106.4982
15.1632 | -106.5231
15.1669 | -106.5423
15.1706 | -106.5596
15.1743 | -106.5789
15.1779 | -106.5975
15.1816 | -106.6129
15.1853 | -106.6276
15.1889 | -106.6381
15.1926 | -106.6443
15.1963 | -106.649
15.2 | -106.6506
15.2036 | -106.648
15.2073 | -106.6378
15.211 | -106.6216
15.2147 | -106.5994
15.2183 | -106.5784
15.222 | -106.5565
15.2257 | -106.5285
15.2293 | -106.4997
15.233 | -106.4696
15.2367 | -106.4376
15.2404 | -106.404
15.244 | -106.3692
15.2477 | -106.3341
15.2514 | -106.2993
15.2551 | -106.266
15.2587 | -106.2345
15.2624 | -106.2043
15.2661 | -106.1759
15.2698 | -106.1517
15.2734 | -106.1303
15.2771 | -106.1106
15.2808 | -106.0976
15.2844 | -106.0861
15.2881 | -106.0765
15.2918 | -106.0693
15.2955 | -106.0641
15.2991 | -106.0597
15.3028 | -106.0592
15.3065 | -106.0633
15.3102 | -106.0687
15.3138 | -106.077
15.3175 | -106.0859
15.3212 | -106.0981
15.3249 | -106.1141
15.3285 | -106.1313
15.3322 | -106.1501
15.3359 | -106.1693
15.3395 | -106.1912
15.3432 | -106.2114
15.3469 | -106.2321
15.3506 | -106.252
15.3542 | -106.2726
15.3579 | -106.2942
15.3616 | -106.3139
15.3653 | -106.3308
15.3689 | -106.3548
15.3726 | -106.3766
15.3763 | -106.3981
15.38 | -106.4139
15.3836 | -106.4323
15.3873 | -106.4526
15.391 | -106.472
15.3946 | -106.4819
15.3983 | -106.4865
15.402 | -106.4919
15.4057 | -106.498
15.4093 | -106.5035
15.413 | -106.5082
15.4167 | -106.5112
15.4204 | -106.5166
15.424 | -106.5228
15.4277 | -106.5263
15.4314 | -106.5282
15.4351 | -106.5302
15.4387 | -106.5296
15.4424 | -106.5245
15.4461 | -106.5211
15.4497 | -106.5161
15.4534 | -106.5116
15.4571 | -106.5069
15.4608 | -106.5025
15.4644 | -106.4975
15.4681 | -106.4899
15.4718 | -106.4832
15.4755 | -106.4791
15.4791 | -106.4751
15.4828 | -106.4705
15.4865 | -106.4661
15.4902 | -106.4603
15.4938 | -106.4546
15.4975 | -106.4507
15.5012 | -106.4468
15.5048 | -106.444
15.5085 | -106.4425
15.5122 | -106.4414
15.5159 | -106.4423
15.5195 | -106.4432
15.5232 | -106.4424
15.5269 | -106.4409
15.5306 | -106.4406
15.5342 | -106.4381
15.5379 | -106.4383
15.5416 | -106.4382
15.5452 | -106.4336
15.5489 | -106.4296
15.5526 | -106.4263
15.5563 | -106.4252
15.5599 | -106.4266
15.5636 | -106.4255
15.5673 | -106.4244
15.571 | -106.4249
15.5746 | -106.4265
15.5783 | -106.4308
15.582 | -106.4377
15.5857 | -106.4427
15.5893 | -106.4445
15.593 | -106.4475
15.5967 | -106.4534
15.6003 | -106.4618
15.604 | -106.4724
15.6077 | -106.479
15.6114 | -106.4867
15.615 | -106.4928
15.6187 | -106.5007
15.6224 | -106.5139
15.6261 | -106.5251
15.6297 | -106.5294
15.6334 | -106.5313
15.6371 | -106.533
15.6408 | -106.5355
15.6444 | -106.5345
15.6481 | -106.5356
15.6518 | -106.5381
15.6554 | -106.5367
15.6591 | -106.5359
15.6628 | -106.5378
15.6665 | -106.5359
15.6701 | -106.5325
15.6738 | -106.5314
15.6775 | -106.5291
15.6812 | -106.5247
15.6848 | -106.5164
15.6885 | -106.5112
15.6922 | -106.5093
15.6959 | -106.504
15.6995 | -106.496
15.7032 | -106.4878
15.7069 | -106.4775
15.7105 | -106.4694
15.7142 | -106.4641
15.7179 | -106.456
15.7216 | -106.4453
15.7252 | -106.4324
15.7289 | -106.4214
15.7326 | -106.4089
15.7363 | -106.3947
15.7399 | -106.3781
15.7436 | -106.3668
15.7473 | -106.3572
15.751 | -106.3462
15.7546 | -106.3351
15.7583 | -106.3248
15.762 | -106.3152
15.7656 | -106.305
15.7693 | -106.2987
15.773 | -106.2909
15.7767 | -106.2821
15.7803 | -106.2746
15.784 | -106.2623
15.7877 | -106.2537
15.7914 | -106.2463
15.795 | -106.238
15.7987 | -106.2284
15.8024 | -106.2167
15.8061 | -106.2058
15.8097 | -106.198
15.8134 | -106.1935
15.8171 | -106.1904
15.8207 | -106.1872
15.8244 | -106.183
15.8281 | -106.1811
15.8318 | -106.1806
15.8354 | -106.1792
15.8391 | -106.1804
15.8428 | -106.1836
15.8465 | -106.1848
15.8501 | -106.1888
15.8538 | -106.1972
15.8575 | -106.2057
15.8612 | -106.213
15.8648 | -106.219
15.8685 | -106.2231
15.8722 | -106.2264
15.8758 | -106.2304
15.8795 | -106.2392
15.8832 | -106.246
15.8869 | -106.2488
15.8905 | -106.2549
15.8942 | -106.2617
15.8979 | -106.2596
15.9016 | -106.2607
15.9052 | -106.2625
15.9089 | -106.264
15.9126 | -106.2647
15.9162 | -106.2632
15.9199 | -106.2588
15.9236 | -106.2534
15.9273 | -106.2452
15.9309 | -106.2361
15.9346 | -106.231
15.9383 | -106.2238
15.942 | -106.2146
15.9456 | -106.2064
15.9493 | -106.1979
15.953 | -106.1909
15.9567 | -106.1825
15.9603 | -106.1742
15.964 | -106.1685
15.9677 | -106.162
15.9713 | -106.1524
15.975 | -106.1434
15.9787 | -106.1374
15.9824 | -106.1324
15.986 | -106.128
15.9897 | -106.1241
15.9934 | -106.1238
15.9971 | -106.1249
16.0007 | -106.1275
16.0044 | -106.1302
16.0081 | -106.1337
16.0118 | -106.1389
16.0154 | -106.1437
16.0191 | -106.1523
16.0228 | -106.1617
16.0264 | -106.1722
16.0301 | -106.1826
16.0338 | -106.1908
16.0375 | -106.1984
16.0411 | -106.2091
16.0448 | -106.2164
16.0485 | -106.2248
16.0522 | -106.2293
16.0558 | -106.2346
16.0595 | -106.2406
16.0632 | -106.248
16.0669 | -106.2547
16.0705 | -106.2605
16.0742 | -106.2651
16.0779 | -106.2668
16.0815 | -106.2654
16.0852 | -106.2629
16.0889 | -106.2605
16.0926 | -106.2573
16.0962 | -106.2543
16.0999 | -106.2495
16.1036 | -106.2425
16.1073 | -106.2336
16.1109 | -106.2245
16.1146 | -106.2146
16.1183 | -106.2066
16.122 | -106.1937
16.1256 | -106.1836
16.1293 | -106.1737
16.133 | -106.1612
16.1366 | -106.15
16.1403 | -106.1411
16.144 | -106.1356
16.1477 | -106.1291
16.1513 | -106.1199
16.155 | -106.1138
16.1587 | -106.1108
16.1624 | -106.1115
16.166 | -106.1112
16.1697 | -106.109
16.1734 | -106.1108
16.1771 | -106.1138
16.1807 | -106.1179
16.1844 | -106.1243
16.1881 | -106.1279
16.1917 | -106.1334
16.1954 | -106.1396
16.1991 | -106.1499
16.2028 | -106.159
16.2064 | -106.1658
16.2101 | -106.174
16.2138 | -106.182
16.2175 | -106.1899
16.2211 | -106.1969
16.2248 | -106.2027
16.2285 | -106.2078
16.2322 | -106.2131
16.2358 | -106.2182
16.2395 | -106.2224
16.2432 | -106.2272
16.2468 | -106.231
16.2505 | -106.2312
16.2542 | -106.2309
16.2579 | -106.2297
16.2615 | -106.2298
16.2652 | -106.2318
16.2689 | -106.232
16.2726 | -106.2284
16.2762 | -106.2276
16.2799 | -106.2248
16.2836 | -106.2282
16.2872 | -106.226
16.2909 | -106.2256
16.2946 | -106.2232
16.2983 | -106.2163
16.3019 | -106.212
16.3056 | -106.2099
16.3093 | -106.2081
16.313 | -106.2064
16.3166 | -106.2065
16.3203 | -106.2054
16.324 | -106.2041
16.3277 | -106.2046
16.3313 | -106.2043
16.335 | -106.204
16.3387 | -106.2029
16.3423 | -106.2018
16.346 | -106.2006
16.3497 | -106.197
16.3534 | -106.1981
16.357 | -106.1989
16.3607 | -106.2005
16.3644 | -106.2022
16.3681 | -106.2013
16.3717 | -106.2029
16.3754 | -106.204
16.3791 | -106.2028
16.3828 | -106.2037
16.3864 | -106.2057
16.3901 | -106.2033
16.3938 | -106.2006
16.3974 | -106.1993
16.4011 | -106.1981
16.4048 | -106.1949
16.4085 | -106.1922
16.4121 | -106.1909
16.4158 | -106.1901
16.4195 | -106.191
16.4232 | -106.1894
16.4268 | -106.1876
16.4305 | -106.1866
16.4342 | -106.1843
16.4379 | -106.1815
16.4415 | -106.1795
16.4452 | -106.1779
16.4489 | -106.1795
16.4525 | -106.1809
16.4562 | -106.1802
16.4599 | -106.1787
16.4636 | -106.1794
16.4672 | -106.1809
16.4709 | -106.182
16.4746 | -106.1803
16.4783 | -106.1828
16.4819 | -106.1871
16.4856 | -106.1896
16.4893 | -106.1956
16.493 | -106.2013
16.4966 | -106.209
16.5003 | -106.2193
16.504 | -106.2294
16.5076 | -106.2374
16.5113 | -106.2467
16.515 | -106.2585
16.5187 | -106.2702
16.5223 | -106.2854
16.526 | -106.2993
16.5297 | -106.314
16.5334 | -106.3309
16.537 | -106.3474
16.5407 | -106.3623
16.5444 | -106.3765
16.5481 | -106.3921
16.5517 | -106.4063
16.5554 | -106.4195
16.5591 | -106.4316
16.5627 | -106.4427
16.5664 | -106.4542
16.5701 | -106.4655
16.5738 | -106.4755
16.5774 | -106.4857
16.5811 | -106.4958
16.5848 | -106.5061
16.5885 | -106.5135
16.5921 | -106.5209
16.5958 | -106.5259
16.5995 | -106.5337
16.6031 | -106.541
16.6068 | -106.5462
16.6105 | -106.5513
16.6142 | -106.5578
16.6178 | -106.5645
16.6215 | -106.567
16.6252 | -106.5748
16.6289 | -106.5814
16.6325 | -106.5861
16.6362 | -106.5924
16.6399 | -106.5987
16.6436 | -106.6022
16.6472 | -106.6065
16.6509 | -106.6096
16.6546 | -106.6141
16.6582 | -106.6195
16.6619 | -106.6209
16.6656 | -106.6242
16.6693 | -106.6282
16.6729 | -106.6315
16.6766 | -106.6359
16.6803 | -106.6399
16.684 | -106.6388
16.6876 | -106.6378
16.6913 | -106.6398
16.695 | -106.6424
16.6987 | -106.6447
16.7023 | -106.6441
16.706 | -106.6422
16.7097 | -106.6425
16.7133 | -106.6449
16.717 | -106.6426
16.7207 | -106.6386
16.7244 | -106.6364
16.728 | -106.6305
16.7317 | -106.6263
16.7354 | -106.6216
16.7391 | -106.6122
16.7427 | -106.6049
16.7464 | -106.596
16.7501 | -106.5877
16.7538 | -106.575
16.7574 | -106.5619
16.7611 | -106.5515
16.7648 | -106.5395
16.7684 | -106.5293
16.7721 | -106.5202
16.7758 | -106.5074
16.7795 | -106.498
16.7831 | -106.487
16.7868 | -106.4748
16.7905 | -106.4629
16.7942 | -106.4513
16.7978 | -106.4387
16.8015 | -106.4264
16.8052 | -106.4143
16.8089 | -106.403
16.8125 | -106.3936
16.8162 | -106.385
16.8199 | -106.3765
16.8235 | -106.367
16.8272 | -106.3579
16.8309 | -106.3484
16.8346 | -106.338
16.8382 | -106.3306
16.8419 | -106.325
16.8456 | -106.32
16.8493 | -106.3161
16.8529 | -106.3124
16.8566 | -106.3116
16.8603 | -106.3114
16.864 | -106.3106
16.8676 | -106.3122
16.8713 | -106.3169
16.875 | -106.321
16.8786 | -106.3246
16.8823 | -106.3315
16.886 | -106.3382
16.8897 | -106.3468
16.8933 | -106.3564
16.897 | -106.3677
16.9007 | -106.3824
16.9044 | -106.3941
16.908 | -106.4037
16.9117 | -106.4136
16.9154 | -106.425
16.919 | -106.4372
16.9227 | -106.4503
16.9264 | -106.4596
16.9301 | -106.4663
16.9337 | -106.4734
16.9374 | -106.4769
16.9411 | -106.4811
16.9448 | -106.486
16.9484 | -106.4888
16.9521 | -106.49
16.9558 | -106.4824
16.9595 | -106.4803
16.9631 | -106.479
16.9668 | -106.4769
16.9705 | -106.4746
16.9741 | -106.4714
16.9778 | -106.4665
16.9815 | -106.4616
16.9852 | -106.4532
16.9888 | -106.4496
16.9925 | -106.4449
16.9962 | -106.4398
16.9999 | -106.43
17.0035 | -106.4171
17.0072 | -106.4069
17.0109 | -106.3984
17.0146 | -106.391
17.0182 | -106.3842
17.0219 | -106.3748
17.0256 | -106.3693
17.0292 | -106.3669
17.0329 | -106.3623
17.0366 | -106.3603
17.0403 | -106.361
17.0439 | -106.3626
17.0476 | -106.3638
17.0513 | -106.3663
17.055 | -106.3689
17.0586 | -106.3728
17.0623 | -106.3792
17.066 | -106.389
17.0697 | -106.4013
17.0733 | -106.4134
17.077 | -106.4276
17.0807 | -106.4402
17.0843 | -106.453
17.088 | -106.4646
17.0917 | -106.4793
17.0954 | -106.4965
17.099 | -106.5111
17.1027 | -106.5271
17.1064 | -106.543
17.1101 | -106.5559
17.1137 | -106.5675
17.1174 | -106.5807
17.1211 | -106.5932
17.1248 | -106.6034
17.1284 | -106.6139
17.1321 | -106.6236
17.1358 | -106.6325
17.1394 | -106.6411
17.1431 | -106.6498
17.1468 | -106.6587
17.1505 | -106.667
17.1541 | -106.6737
17.1578 | -106.6765
17.1615 | -106.6796
17.1652 | -106.6834
17.1688 | -106.6864
17.1725 | -106.6907
17.1762 | -106.6919
17.1799 | -106.6926
17.1835 | -106.6944
17.1872 | -106.6947
17.1909 | -106.6928
17.1945 | -106.6923
17.1982 | -106.6927
17.2019 | -106.6917
17.2056 | -106.6911
17.2092 | -106.6895
17.2129 | -106.6881
17.2166 | -106.686
17.2203 | -106.6828
17.2239 | -106.6808
17.2276 | -106.6807
17.2313 | -106.6798
17.235 | -106.677
17.2386 | -106.6714
17.2423 | -106.6679
17.246 | -106.6664
17.2496 | -106.6615
17.2533 | -106.6537
17.257 | -106.6473
17.2607 | -106.6398
17.2643 | -106.6328
17.268 | -106.6242
17.2717 | -106.6135
17.2754 | -106.6005
17.279 | -106.5877
17.2827 | -106.5743
17.2864 | -106.5574
17.2901 | -106.5394
17.2937 | -106.523
17.2974 | -106.5081
17.3011 | -106.4952
17.3047 | -106.4793
17.3084 | -106.4599
17.3121 | -106.4386
17.3158 | -106.4167
17.3194 | -106.3967
17.3231 | -106.3755
17.3268 | -106.3558
17.3305 | -106.336
17.3341 | -106.317
17.3378 | -106.2958
17.3415 | -106.2759
17.3451 | -106.2571
17.3488 | -106.2402
17.3525 | -106.2228
17.3562 | -106.2073
17.3598 | -106.1939
17.3635 | -106.1813
17.3672 | -106.1725
17.3709 | -106.1655
17.3745 | -106.159
17.3782 | -106.1524
17.3819 | -106.146
17.3856 | -106.1411
17.3892 | -106.1375
17.3929 | -106.1373
17.3966 | -106.1409
17.4002 | -106.1466
17.4039 | -106.1542
17.4076 | -106.1615
17.4113 | -106.1703
17.4149 | -106.1796
17.4186 | -106.1877
17.4223 | -106.2012
17.426 | -106.2168
17.4296 | -106.2258
17.4333 | -106.235
17.437 | -106.2452
17.4407 | -106.2567
17.4443 | -106.2678
17.448 | -106.2787
17.4517 | -106.2867
17.4553 | -106.2941
17.459 | -106.3008
17.4627 | -106.3072
17.4664 | -106.3129
17.47 | -106.3168
17.4737 | -106.3187
17.4774 | -106.3187
17.4811 | -106.3154
17.4847 | -106.3093
17.4884 | -106.3108
17.4921 | -106.3126
17.4958 | -106.3129
17.4994 | -106.3119
17.5031 | -106.3118
17.5068 | -106.3083
17.5104 | -106.3049
17.5141 | -106.3023
17.5178 | -106.3017
17.5215 | -106.3023
17.5251 | -106.3033
17.5288 | -106.3066
17.5325 | -106.3096
17.5362 | -106.3139
17.5398 | -106.319
17.5435 | -106.3241
17.5472 | -106.327
17.5509 | -106.3294
17.5545 | -106.3356
17.5582 | -106.3451
17.5619 | -106.3556
17.5655 | -106.3677
17.5692 | -106.3803
17.5729 | -106.3944
17.5766 | -106.4081
17.5802 | -106.4224
17.5839 | -106.4383
17.5876 | -106.4572
17.5913 | -106.4751
17.5949 | -106.4934
17.5986 | -106.5111
17.6023 | -106.5287
17.606 | -106.5491
17.6096 | -106.5696
17.6133 | -106.5881
17.617 | -106.6023
17.6206 | -106.6153
17.6243 | -106.6293
17.628 | -106.6429
17.6317 | -106.6565
17.6353 | -106.6706
17.639 | -106.6813
17.6427 | -106.6884
17.6464 | -106.6939
17.65 | -106.6992
17.6537 | -106.7043
17.6574 | -106.7056
17.661 | -106.7044
17.6647 | -106.7069
17.6684 | -106.7074
17.6721 | -106.7034
17.6757 | -106.6952
17.6794 | -106.6891
17.6831 | -106.6841
17.6868 | -106.679
17.6904 | -106.6693
17.6941 | -106.6594
17.6978 | -106.6479
17.7015 | -106.6403
17.7051 | -106.6323
17.7088 | -106.6223
17.7125 | -106.6127
17.7161 | -106.6055
17.7198 | -106.5976
17.7235 | -106.5884
17.7272 | -106.5833
17.7308 | -106.5794
17.7345 | -106.5724
17.7382 | -106.5641
17.7419 | -106.558
17.7455 | -106.5522
17.7492 | -106.5462
17.7529 | -106.544
17.7566 | -106.5416
17.7602 | -106.54
17.7639 | -106.5418
17.7676 | -106.543
17.7712 | -106.5441
17.7749 | -106.5457
17.7786 | -106.5484
17.7823 | -106.5516
17.7859 | -106.5513
17.7896 | -106.5515
17.7933 | -106.5537
17.797 | -106.5573
17.8006 | -106.5619
17.8043 | -106.5652
17.808 | -106.5679
17.8117 | -106.5694
17.8153 | -106.5695
17.819 | -106.5708
17.8227 | -106.5693
17.8263 | -106.5676
17.83 | -106.5643
17.8337 | -106.5624
17.8374 | -106.5576
17.841 | -106.5493
17.8447 | -106.5413
17.8484 | -106.5343
17.8521 | -106.5263
17.8557 | -106.5205
17.8594 | -106.5098
17.8631 | -106.4993
17.8668 | -106.489
17.8704 | -106.4808
17.8741 | -106.4759
17.8778 | -106.4674
17.8814 | -106.4576
17.8851 | -106.4484
17.8888 | -106.4377
17.8925 | -106.4302
17.8961 | -106.4244
17.8998 | -106.4174
17.9035 | -106.4101
17.9072 | -106.4038
17.9108 | -106.3994
17.9145 | -106.3959
17.9182 | -106.3911
17.9219 | -106.3903
17.9255 | -106.3918
17.9292 | -106.3964
17.9329 | -106.4004
17.9365 | -106.4027
17.9402 | -106.405
17.9439 | -106.4091
17.9476 | -106.4145
17.9512 | -106.4213
17.9549 | -106.4253
17.9586 | -106.4308
17.9623 | -106.4384
17.9659 | -106.4463
17.9696 | -106.454
17.9733 | -106.4637
17.9769 | -106.4761
17.9806 | -106.4856
17.9843 | -106.4934
17.988 | -106.5057
17.9916 | -106.5188
17.9953 | -106.5318
17.999 | -106.5422
18.0027 | -106.5534
18.0063 | -106.5659
18.01 | -106.5811
18.0137 | -106.5939
18.0174 | -106.6024
18.021 | -106.6135
18.0247 | -106.6263
18.0284 | -106.6379
18.032 | -106.6479
18.0357 | -106.656
18.0394 | -106.664
18.0431 | -106.6747
18.0467 | -106.6845
18.0504 | -106.6911
18.0541 | -106.695
18.0578 | -106.6982
18.0614 | -106.701
18.0651 | -106.7028
18.0688 | -106.7047
18.0725 | -106.7063
18.0761 | -106.7059
18.0798 | -106.706
18.0835 | -106.7065
18.0871 | -106.7056
18.0908 | -106.7021
18.0945 | -106.7014
18.0982 | -106.7032
18.1018 | -106.7058
18.1055 | -106.7067
18.1092 | -106.7078
18.1129 | -106.7064
18.1165 | -106.7088
18.1202 | -106.7101
18.1239 | -106.714
18.1276 | -106.7183
18.1312 | -106.7174
18.1349 | -106.7189
18.1386 | -106.7244
18.1422 | -106.7324
18.1459 | -106.7384
18.1496 | -106.7419
18.1533 | -106.745
18.1569 | -106.7474
18.1606 | -106.7507
18.1643 | -106.7528
18.168 | -106.7528
18.1716 | -106.7528
18.1753 | -106.7516
18.179 | -106.7501
18.1827 | -106.747
18.1863 | -106.7421
18.19 | -106.735
18.1937 | -106.7291
18.1973 | -106.7229
18.201 | -106.7156
18.2047 | -106.7089
18.2084 | -106.7014
18.212 | -106.6942
18.2157 | -106.6856
18.2194 | -106.6764
18.2231 | -106.6669
18.2267 | -106.6565
18.2304 | -106.6449
18.2341 | -106.6336
18.2378 | -106.6233
18.2414 | -106.6125
18.2451 | -106.599
18.2488 | -106.5873
18.2524 | -106.5764
18.2561 | -106.5689
18.2598 | -106.5623
18.2635 | -106.554
18.2671 | -106.5452
18.2708 | -106.5403
18.2745 | -106.537
18.2782 | -106.5327
18.2818 | -106.5267
18.2855 | -106.523
18.2892 | -106.5236
18.2929 | -106.5247
18.2965 | -106.528
18.3002 | -106.5318
18.3039 | -106.5375
18.3075 | -106.542
18.3112 | -106.548
18.3149 | -106.5567
18.3186 | -106.5651
18.3222 | -106.5763
18.3259 | -106.5887
18.3296 | -106.6004
18.3333 | -106.6117
18.3369 | -106.6218
18.3406 | -106.6309
18.3443 | -106.6415
18.348 | -106.6533
18.3516 | -106.6652
18.3553 | -106.675
18.359 | -106.6838
18.3626 | -106.6909
18.3663 | -106.699
18.37 | -106.7072
18.3737 | -106.7147
18.3773 | -106.7202
18.381 | -106.727
18.3847 | -106.7301
18.3884 | -106.7325
18.392 | -106.7356
18.3957 | -106.7375
18.3994 | -106.7372
18.403 | -106.7353
18.4067 | -106.7346
18.4104 | -106.7335
18.4141 | -106.7349
18.4177 | -106.7357
18.4214 | -106.7343
18.4251 | -106.733
18.4288 | -106.7335
18.4324 | -106.7308
18.4361 | -106.7288
18.4398 | -106.7268
18.4435 | -106.7263
18.4471 | -106.7282
18.4508 | -106.7304
18.4545 | -106.7297
18.4581 | -106.7309
18.4618 | -106.7332
18.4655 | -106.7346
18.4692 | -106.7354
18.4728 | -106.7364
18.4765 | -106.7369
18.4802 | -106.7363
18.4839 | -106.7359
18.4875 | -106.7362
18.4912 | -106.7363
18.4949 | -106.7376
18.4986 | -106.7396
18.5022 | -106.7437
18.5059 | -106.745
18.5096 | -106.7453
18.5132 | -106.7472
18.5169 | -106.7496
18.5206 | -106.7493
18.5243 | -106.7494
18.5279 | -106.7499
18.5316 | -106.7505
18.5353 | -106.7502
18.539 | -106.7502
18.5426 | -106.7542
18.5463 | -106.7537
18.55 | -106.7529
18.5537 | -106.7521
18.5573 | -106.7528
18.561 | -106.7523
18.5647 | -106.7519
18.5683 | -106.7504
18.572 | -106.751
18.5757 | -106.748
18.5794 | -106.7449
18.583 | -106.7462
18.5867 | -106.746
18.5904 | -106.7433
18.5941 | -106.7413
18.5977 | -106.7389
18.6014 | -106.736
18.6051 | -106.7344
18.6088 | -106.7337
18.6124 | -106.731
18.6161 | -106.7278
18.6198 | -106.7247
18.6234 | -106.7194
18.6271 | -106.7138
18.6308 | -106.7091
18.6345 | -106.7045
18.6381 | -106.7013
18.6418 | -106.6976
18.6455 | -106.6934
18.6492 | -106.6915
18.6528 | -106.6882
18.6565 | -106.6858
18.6602 | -106.6822
18.6639 | -106.678
18.6675 | -106.6733
18.6712 | -106.6709
18.6749 | -106.6707
18.6785 | -106.6658
18.6822 | -106.6613
18.6859 | -106.6567
18.6896 | -106.6514
18.6932 | -106.6457
18.6969 | -106.6422
18.7006 | -106.6387
18.7043 | -106.635
18.7079 | -106.632
18.7116 | -106.6282
18.7153 | -106.6259
18.7189 | -106.625
18.7226 | -106.6216
18.7263 | -106.6188
18.73 | -106.6153
18.7336 | -106.6149
18.7373 | -106.6149
18.741 | -106.6135
18.7447 | -106.6148
18.7483 | -106.6169
18.752 | -106.6186
18.7557 | -106.6225
18.7594 | -106.6267
18.763 | -106.6312
18.7667 | -106.6338
18.7704 | -106.6402
18.774 | -106.6464
18.7777 | -106.6518
18.7814 | -106.6575
18.7851 | -106.6688
18.7887 | -106.6788
18.7924 | -106.6866
18.7961 | -106.6952
18.7998 | -106.7041
18.8034 | -106.715
18.8071 | -106.7255
18.8108 | -106.7375
18.8145 | -106.7514
18.8181 | -106.7647
18.8218 | -106.7756
18.8255 | -106.7864
18.8291 | -106.7965
18.8328 | -106.8051
18.8365 | -106.8149
18.8402 | -106.8265
18.8438 | -106.839
18.8475 | -106.8495
18.8512 | -106.8582
18.8549 | -106.8663
18.8585 | -106.8742
18.8622 | -106.8826
18.8659 | -106.8924
18.8696 | -106.8985
18.8732 | -106.9043
18.8769 | -106.9076
18.8806 | -106.913
18.8842 | -106.9165
18.8879 | -106.918
18.8916 | -106.9196
18.8953 | -106.9198
18.8989 | -106.9158
18.9026 | -106.912
18.9063 | -106.9074
18.91 | -106.8985
18.9136 | -106.8911
18.9173 | -106.8847
18.921 | -106.8774
18.9247 | -106.8678
18.9283 | -106.8582
18.932 | -106.8471
18.9357 | -106.8364
18.9393 | -106.8233
18.943 | -106.8073
18.9467 | -106.7959
18.9504 | -106.784
18.954 | -106.7703
18.9577 | -106.7567
18.9614 | -106.744
18.9651 | -106.7307
18.9687 | -106.7154
18.9724 | -106.7
18.9761 | -106.6837
18.9798 | -106.6671
18.9834 | -106.6511
18.9871 | -106.6352
18.9908 | -106.621
18.9944 | -106.6066
18.9981 | -106.5938
19.0018 | -106.5811
19.0055 | -106.5696
19.0091 | -106.5607
19.0128 | -106.5508
19.0165 | -106.5391
19.0202 | -106.5285
19.0238 | -106.5178
19.0275 | -106.5048
19.0312 | -106.495
19.0348 | -106.4891
19.0385 | -106.4838
19.0422 | -106.4782
19.0459 | -106.474
19.0495 | -106.4714
19.0532 | -106.4683
19.0569 | -106.4664
19.0606 | -106.4681
19.0642 | -106.4706
19.0679 | -106.4766
19.0716 | -106.4806
19.0753 | -106.491
19.0789 | -106.4947
19.0826 | -106.4971
19.0863 | -106.502
19.0899 | -106.5087
19.0936 | -106.5175
19.0973 | -106.524
19.101 | -106.5308
19.1046 | -106.5376
19.1083 | -106.5436
19.112 | -106.5503
19.1157 | -106.5581
19.1193 | -106.5647
19.123 | -106.5698
19.1267 | -106.5765
19.1304 | -106.5822
19.134 | -106.587
19.1377 | -106.5907
19.1414 | -106.5977
19.145 | -106.6026
19.1487 | -106.6057
19.1524 | -106.6049
19.1561 | -106.6051
19.1597 | -106.6066
19.1634 | -106.6071
19.1671 | -106.607
19.1708 | -106.6069
19.1744 | -106.6046
19.1781 | -106.6002
19.1818 | -106.5985
19.1855 | -106.5939
19.1891 | -106.5878
19.1928 | -106.5815
19.1965 | -106.5759
19.2001 | -106.5713
19.2038 | -106.5671
19.2075 | -106.5579
19.2112 | -106.5502
19.2148 | -106.5442
19.2185 | -106.5359
19.2222 | -106.5294
19.2259 | -106.5217
19.2295 | -106.5107
19.2332 | -106.4987
19.2369 | -106.4859
19.2406 | -106.4734
19.2442 | -106.4618
19.2479 | -106.4522
19.2516 | -106.4432
19.2552 | -106.4339
19.2589 | -106.4247
19.2626 | -106.4146
19.2663 | -106.4051
19.2699 | -106.3944
19.2736 | -106.3849
19.2773 | -106.3751
19.281 | -106.3639
19.2846 | -106.3533
19.2883 | -106.345
19.292 | -106.336
19.2957 | -106.3271
19.2993 | -106.3183
19.303 | -106.3087
19.3067 | -106.2995
19.3103 | -106.2897
19.314 | -106.2814
19.3177 | -106.2738
19.3214 | -106.2682
19.325 | -106.2629
19.3287 | -106.2584
19.3324 | -106.2533
19.3361 | -106.2498
19.3397 | -106.2468
19.3434 | -106.2464
19.3471 | -106.2475
19.3508 | -106.2486
19.3544 | -106.2522
19.3581 | -106.2567
19.3618 | -106.2633
19.3654 | -106.2721
19.3691 | -106.2818
19.3728 | -106.2933
19.3765 | -106.3083
19.3801 | -106.3247
19.3838 | -106.3419
19.3875 | -106.3579
19.3912 | -106.3739
19.3948 | -106.3923
19.3985 | -106.4107
19.4022 | -106.4306
19.4059 | -106.4532
19.4095 | -106.476
19.4132 | -106.4984
19.4169 | -106.5199
19.4205 | -106.5414
19.4242 | -106.5634
19.4279 | -106.5894
19.4316 | -106.6126
19.4352 | -106.6324
19.4389 | -106.6514
19.4426 | -106.6712
19.4463 | -106.6903
19.4499 | -106.7081
19.4536 | -106.726
19.4573 | -106.7392
19.4609 | -106.7508
19.4646 | -106.7596
19.4683 | -106.7675
19.472 | -106.7734
19.4756 | -106.7777
19.4793 | -106.7787
19.483 | -106.7797
19.4867 | -106.7778
19.4903 | -106.7733
19.494 | -106.7652
19.4977 | -106.7536
19.5014 | -106.7443
19.505 | -106.7365
19.5087 | -106.7246
19.5124 | -106.7116
19.516 | -106.694
19.5197 | -106.6734
19.5234 | -106.6538
19.5271 | -106.6344
19.5307 | -106.6129
19.5344 | -106.5916
19.5381 | -106.5701
19.5418 | -106.5499
19.5454 | -106.5294
19.5491 | -106.5089
19.5528 | -106.4861
19.5565 | -106.4644
19.5601 | -106.4441
19.5638 | -106.4244
19.5675 | -106.4082
19.5711 | -106.3937
19.5748 | -106.3822
19.5785 | -106.3683
19.5822 | -106.3551
19.5858 | -106.3443
19.5895 | -106.3353
19.5932 | -106.3275
19.5969 | -106.3193
19.6005 | -106.3133
19.6042 | -106.3203
19.6079 | -106.3167
19.6116 | -106.3156
19.6152 | -106.3152
19.6189 | -106.3156
19.6226 | -106.3181
19.6262 | -106.3218
19.6299 | -106.327
19.6336 | -106.333
19.6373 | -106.3419
19.6409 | -106.3497
19.6446 | -106.3576
19.6483 | -106.3665
19.652 | -106.3781
19.6556 | -106.3887
19.6593 | -106.3997
19.663 | -106.411
19.6667 | -106.4233
19.6703 | -106.437
19.674 | -106.4505
19.6777 | -106.4646
19.6813 | -106.478
19.685 | -106.4905
19.6887 | -106.5012
19.6924 | -106.512
19.696 | -106.5231
19.6997 | -106.535
19.7034 | -106.5458
19.7071 | -106.5532
19.7107 | -106.5603
19.7144 | -106.5692
19.7181 | -106.5755
19.7218 | -106.5792
19.7254 | -106.5821
19.7291 | -106.5851
19.7328 | -106.5854
19.7364 | -106.5874
19.7401 | -106.5863
19.7438 | -106.5836
19.7475 | -106.5804
19.7511 | -106.5776
19.7548 | -106.5722
19.7585 | -106.5666
19.7622 | -106.5599
19.7658 | -106.5529
19.7695 | -106.5441
19.7732 | -106.5368
19.7768 | -106.5298
19.7805 | -106.5201
19.7842 | -106.5082
19.7879 | -106.498
19.7915 | -106.4879
19.7952 | -106.4806
19.7989 | -106.4738
19.8026 | -106.4644
19.8062 | -106.4544
19.8099 | -106.4472
19.8136 | -106.4407
19.8173 | -106.4352
19.8209 | -106.4302
19.8246 | -106.4281
19.8283 | -106.4262
19.8319 | -106.4271
19.8356 | -106.4278
19.8393 | -106.428
19.843 | -106.4284
19.8466 | -106.4287
19.8503 | -106.4307
19.854 | -106.4342
19.8577 | -106.4409
19.8613 | -106.4475
19.865 | -106.4536
19.8687 | -106.4623
19.8724 | -106.4702
19.876 | -106.4788
19.8797 | -106.4883
19.8834 | -106.4992
19.887 | -106.5099
19.8907 | -106.5214
19.8944 | -106.5311
19.8981 | -106.5387
19.9017 | -106.5475
19.9054 | -106.5558
19.9091 | -106.5636
19.9128 | -106.5696
19.9164 | -106.578
19.9201 | -106.5865
19.9238 | -106.5948
19.9275 | -106.6015
19.9311 | -106.6064
19.9348 | -106.611
19.9385 | -106.6152
19.9421 | -106.6184
19.9458 | -106.6181
19.9495 | -106.6174
19.9532 | -106.616
19.9568 | -106.6124
19.9605 | -106.6099
19.9642 | -106.6064
19.9679 | -106.603
19.9715 | -106.6008
19.9752 | -106.5969
19.9789 | -106.5897
19.9826 | -106.5801
19.9862 | -106.5707
19.9899 | -106.5629
19.9936 | -106.5567
19.9972 | -106.5489
20.0009 | -106.5425
20.0046 | -106.5361
20.0083 | -106.5287
20.0119 | -106.5189
20.0156 | -106.5117
20.0193 | -106.504
20.023 | -106.4991
20.0266 | -106.494
20.0303 | -106.4882
20.034 | -106.4812
20.0377 | -106.4745
20.0413 | -106.4696
20.045 | -106.4662
20.0487 | -106.4606
20.0523 | -106.4558
20.056 | -106.4497
20.0597 | -106.4466
20.0634 | -106.4432
20.067 | -106.4392
20.0707 | -106.4349
20.0744 | -106.4294
20.0781 | -106.4225
20.0817 | -106.417
20.0854 | -106.4116
20.0891 | -106.4064
20.0927 | -106.3983
20.0964 | -106.3893
20.1001 | -106.3831
20.1038 | -106.3764
20.1074 | -106.3695
20.1111 | -106.3625
20.1148 | -106.3555
20.1185 | -106.3489
20.1221 | -106.3442
20.1258 | -106.339
20.1295 | -106.3308
20.1332 | -106.3231
20.1368 | -106.3132
20.1405 | -106.3028
20.1442 | -106.293
20.1478 | -106.2834
20.1515 | -106.2764
20.1552 | -106.2678
20.1589 | -106.2598
20.1625 | -106.2522
20.1662 | -106.2455
20.1699 | -106.2383
20.1736 | -106.2321
20.1772 | -106.2263
20.1809 | -106.2231
20.1846 | -106.2218
20.1883 | -106.2179
20.1919 | -106.2145
20.1956 | -106.2117
20.1993 | -106.2083
20.2029 | -106.2044
20.2066 | -106.2009
20.2103 | -106.1973
20.214 | -106.1954
20.2176 | -106.1962
20.2213 | -106.1954
20.225 | -106.1952
20.2287 | -106.1954
20.2323 | -106.1975
20.236 | -106.2003
20.2397 | -106.2029
20.2434 | -106.2066
20.247 | -106.209
20.2507 | -106.2077
20.2544 | -106.2063
20.258 | -106.2044
20.2617 | -106.2032
20.2654 | -106.2027
20.2691 | -106.2032
20.2727 | -106.2053
20.2764 | -106.2047
20.2801 | -106.2027
20.2838 | -106.2014
20.2874 | -106.1993
20.2911 | -106.1973
20.2948 | -106.1973
20.2985 | -106.1961
20.3021 | -106.1918
20.3058 | -106.1877
20.3095 | -106.1824
20.3131 | -106.179
20.3168 | -106.1731
20.3205 | -106.1665
20.3242 | -106.1622
20.3278 | -106.1573
20.3315 | -106.1514
20.3352 | -106.1457
20.3389 | -106.1395
20.3425 | -106.13
20.3462 | -106.1187
20.3499 | -106.1096
20.3536 | -106.1017
20.3572 | -106.0919
20.3609 | -106.0822
20.3646 | -106.0744
20.3682 | -106.0636
20.3719 | -106.0518
20.3756 | -106.0396
20.3793 | -106.03
20.3829 | -106.019
20.3866 | -106.0059
20.3903 | -105.9947
20.394 | -105.9834
20.3976 | -105.9738
20.4013 | -105.9646
20.405 | -105.9556
20.4087 | -105.9447
20.4123 | -105.9355
20.416 | -105.928
20.4197 | -105.9205
20.4233 | -105.9152
20.427 | -105.9105
20.4307 | -105.9067
20.4344 | -105.9042
20.438 | -105.9025
20.4417 | -105.8977
20.4454 | -105.8938
20.4491 | -105.8936
20.4527 | -105.8947
20.4564 | -105.8953
20.4601 | -105.897
20.4637 | -105.8974
20.4674 | -105.899
20.4711 | -105.9026
20.4748 | -105.9055
20.4784 | -105.9096
20.4821 | -105.9138
20.4858 | -105.9193
20.4895 | -105.9283
20.4931 | -105.9364
20.4968 | -105.9455
20.5005 | -105.956
20.5042 | -105.9681
20.5078 | -105.9798
20.5115 | -105.9911
20.5152 | -106.0029
20.5188 | -106.0167
20.5225 | -106.0312
20.5262 | -106.0438
20.5299 | -106.0561
20.5335 | -106.0684
20.5372 | -106.0809
20.5409 | -106.0953
20.5446 | -106.109
20.5482 | -106.1226
20.5519 | -106.1352
20.5556 | -106.1474
20.5593 | -106.1576
20.5629 | -106.1663
20.5666 | -106.1774
20.5703 | -106.1876
20.5739 | -106.1973
20.5776 | -106.2059
20.5813 | -106.211
20.585 | -106.2156
20.5886 | -106.2247
20.5923 | -106.2327
20.596 | -106.2398
20.5997 | -106.2475
20.6033 | -106.2529
20.607 | -106.2594
20.6107 | -106.2664
20.6144 | -106.2685
20.618 | -106.2701
20.6217 | -106.2707
20.6254 | -106.2714
20.629 | -106.2728
20.6327 | -106.2766
20.6364 | -106.2788
20.6401 | -106.279
20.6437 | -106.2794
20.6474 | -106.2799
20.6511 | -106.2813
20.6548 | -106.2803
20.6584 | -106.2797
20.6621 | -106.2796
20.6658 | -106.2776
20.6695 | -106.2782
20.6731 | -106.2757
20.6768 | -106.2752
20.6805 | -106.2731
20.6841 | -106.2669
20.6878 | -106.2628
20.6915 | -106.2601
20.6952 | -106.2573
20.6988 | -106.2536
20.7025 | -106.2491
20.7062 | -106.245
20.7099 | -106.242
20.7135 | -106.2382
20.7172 | -106.2332
20.7209 | -106.2276
20.7246 | -106.2227
20.7282 | -106.2172
20.7319 | -106.2118
20.7356 | -106.2053
20.7392 | -106.2025
20.7429 | -106.1987
20.7466 | -106.1935
20.7503 | -106.1894
20.7539 | -106.1822
20.7576 | -106.1781
20.7613 | -106.1757
20.765 | -106.1691
20.7686 | -106.1622
20.7723 | -106.1536
20.776 | -106.1483
20.7797 | -106.1437
20.7833 | -106.1369
20.787 | -106.1285
20.7907 | -106.1203
20.7943 | -106.1115
20.798 | -106.1012
20.8017 | -106.0921
20.8054 | -106.0845
20.809 | -106.0739
20.8127 | -106.0606
20.8164 | -106.0496
20.8201 | -106.039
20.8237 | -106.0267
20.8274 | -106.0153
20.8311 | -106.0036
20.8347 | -105.9925
20.8384 | -105.9804
20.8421 | -105.9677
20.8458 | -105.9551
20.8494 | -105.9417
20.8531 | -105.9272
20.8568 | -105.9114
20.8605 | -105.8959
20.8641 | -105.8804
20.8678 | -105.8652
20.8715 | -105.8526
20.8752 | -105.8416
20.8788 | -105.8281
20.8825 | -105.8128
20.8862 | -105.8006
20.8898 | -105.7881
20.8935 | -105.7741
20.8972 | -105.7624
20.9009 | -105.7506
20.9045 | -105.7396
20.9082 | -105.7315
20.9119 | -105.723
20.9156 | -105.718
20.9192 | -105.7136
20.9229 | -105.7089
20.9266 | -105.7048
20.9303 | -105.7018
20.9339 | -105.6987
20.9376 | -105.6947
20.9413 | -105.695
20.9449 | -105.6955
20.9486 | -105.6961
20.9523 | -105.6977
20.956 | -105.7009
20.9596 | -105.7044
20.9633 | -105.7098
20.967 | -105.7139
20.9707 | -105.7201
20.9743 | -105.7248
20.978 | -105.7304
20.9817 | -105.7379
20.9854 | -105.7451
20.989 | -105.7506
20.9927 | -105.7561
20.9964 | -105.763
21 | -105.7718
21.0037 | -105.778
21.0074 | -105.7842
21.0111 | -105.7906
21.0147 | -105.7974
21.0184 | -105.806
21.0221 | -105.8125
21.0258 | -105.8179
21.0294 | -105.8242
21.0331 | -105.8328
21.0368 | -105.8373
21.0405 | -105.8414
21.0441 | -105.8444
21.0478 | -105.8486
21.0515 | -105.8547
21.0551 | -105.8595
21.0588 | -105.863
21.0625 | -105.8661
21.0662 | -105.8692
21.0698 | -105.8741
21.0735 | -105.8749
21.0772 | -105.8739
21.0809 | -105.8748
21.0845 | -105.8781
21.0882 | -105.8806
21.0919 | -105.8837
21.0956 | -105.8882
21.0992 | -105.8914
21.1029 | -105.8941
21.1066 | -105.8995
21.1102 | -105.9061
21.1139 | -105.9098
21.1176 | -105.9134
21.1213 | -105.9188
21.1249 | -105.9225
21.1286 | -105.9275
21.1323 | -105.9329
21.136 | -105.9385
21.1396 | -105.9445
21.1433 | -105.9503
21.147 | -105.9568
21.1506 | -105.9638
21.1543 | -105.9714
21.158 | -105.9793
21.1617 | -105.9867
21.1653 | -105.9928
21.169 | -105.9992
21.1727 | -106.0061
21.1764 | -106.0158
21.18 | -106.0253
21.1837 | -106.0329
21.1874 | -106.0416
21.1911 | -106.0503
21.1947 | -106.0581
21.1984 | -106.0663
21.2021 | -106.0757
21.2057 | -106.0841
21.2094 | -106.0915
21.2131 | -106.0987
21.2168 | -106.1018
21.2204 | -106.1064
21.2241 | -106.1133
21.2278 | -106.1185
21.2315 | -106.1233
21.2351 | -106.1283
21.2388 | -106.1341
21.2425 | -106.1366
21.2462 | -106.1398
21.2498 | -106.1445
21.2535 | -106.1488
21.2572 | -106.1515
21.2608 | -106.1539
21.2645 | -106.1559
21.2682 | -106.1592
21.2719 | -106.1618
21.2755 | -106.165
21.2792 | -106.167
21.2829 | -106.1677
21.2866 | -106.1677
21.2902 | -106.1693
21.2939 | -106.1727
21.2976 | -106.1765
21.3013 | -106.1784
21.3049 | -106.1786
21.3086 | -106.1775
21.3123 | -106.176
21.3159 | -106.1759
21.3196 | -106.1759
21.3233 | -106.1765
21.327 | -106.1785
21.3306 | -106.1769
21.3343 | -106.1752
21.338 | -106.1743
21.3417 | -106.1735
21.3453 | -106.1743
21.349 | -106.1731
21.3527 | -106.1714
21.3564 | -106.17
21.36 | -106.1675
21.3637 | -106.167
21.3674 | -106.1671
21.371 | -106.165
21.3747 | -106.1635
21.3784 | -106.1618
21.3821 | -106.1586
21.3857 | -106.155
21.3894 | -106.1499
21.3931 | -106.1428
21.3968 | -106.1361
21.4004 | -106.13
21.4041 | -106.1247
21.4078 | -106.1184
21.4115 | -106.1125
21.4151 | -106.1079
21.4188 | -106.1034
21.4225 | -106.0985
21.4261 | -106.0932
21.4298 | -106.0881
21.4335 | -106.0829
21.4372 | -106.0782
21.4408 | -106.0736
21.4445 | -106.0684
21.4482 | -106.0641
21.4519 | -106.0602
21.4555 | -106.056
21.4592 | -106.051
21.4629 | -106.0481
21.4666 | -106.0457
21.4702 | -106.0409
21.4739 | -106.0371
21.4776 | -106.034
21.4812 | -106.0322
21.4849 | -106.0296
21.4886 | -106.0267
21.4923 | -106.0254
21.4959 | -106.0263
21.4996 | -106.0269
21.5033 | -106.0291
21.507 | -106.0313
21.5106 | -106.0339
21.5143 | -106.0367
21.518 | -106.0407
21.5216 | -106.0463
21.5253 | -106.0501
21.529 | -106.0526
21.5327 | -106.0568
21.5363 | -106.0617
21.54 | -106.0679
21.5437 | -106.0737
21.5474 | -106.0772
21.551 | -106.0797
21.5547 | -106.0832
21.5584 | -106.0862
21.5621 | -106.0883
21.5657 | -106.0921
21.5694 | -106.0949
21.5731 | -106.098
21.5767 | -106.1009
21.5804 | -106.1046
21.5841 | -106.1091
21.5878 | -106.1124
21.5914 | -106.1142
21.5951 | -106.1158
21.5988 | -106.1161
21.6025 | -106.1155
21.6061 | -106.1149
21.6098 | -106.1176
21.6135 | -106.1185
21.6172 | -106.1182
21.6208 | -106.1167
21.6245 | -106.1132
21.6282 | -106.1116
21.6318 | -106.1097
21.6355 | -106.1096
21.6392 | -106.1078
21.6429 | -106.1038
21.6465 | -106.1005
21.6502 | -106.0972
21.6539 | -106.0939
21.6576 | -106.0916
21.6612 | -106.0894
21.6649 | -106.0871
21.6686 | -106.0842
21.6723 | -106.0793
21.6759 | -106.0768
21.6796 | -106.0733
21.6833 | -106.0677
21.6869 | -106.0664
21.6906 | -106.0652
21.6943 | -106.0657
21.698 | -106.0661
21.7016 | -106.0655
21.7053 | -106.0675
21.709 | -106.0704
21.7127 | -106.0738
21.7163 | -106.0745
21.72 | -106.0731
21.7237 | -106.0745
21.7274 | -106.0827
21.731 | -106.0848
21.7347 | -106.0872
21.7384 | -106.0875
21.742 | -106.0867
21.7457 | -106.0883
21.7494 | -106.0916
21.7531 | -106.0914
21.7567 | -106.0908
21.7604 | -106.0912
21.7641 | -106.0923
21.7678 | -106.0921
21.7714 | -106.0891
21.7751 | -106.0856
21.7788 | -106.0839
21.7825 | -106.0799
21.7861 | -106.0727
21.7898 | -106.0626
21.7935 | -106.0529
21.7971 | -106.0425
21.8008 | -106.0307
21.8045 | -106.0167
21.8082 | -106.0027
21.8118 | -105.9876
21.8155 | -105.9701
21.8192 | -105.9521
21.8229 | -105.9358
21.8265 | -105.9183
21.8302 | -105.9002
21.8339 | -105.8831
21.8376 | -105.8654
21.8412 | -105.8455
21.8449 | -105.8237
21.8486 | -105.8015
21.8522 | -105.7819
21.8559 | -105.7636
21.8596 | -105.7437
21.8633 | -105.722
21.8669 | -105.7027
21.8706 | -105.6853
21.8743 | -105.6662
21.878 | -105.6473
21.8816 | -105.6284
21.8853 | -105.6089
21.889 | -105.588
21.8926 | -105.5693
21.8963 | -105.5534
21.9 | -105.54
21.9037 | -105.5255
21.9073 | -105.5105
21.911 | -105.4967
21.9147 | -105.4852
21.9184 | -105.4733
21.922 | -105.4636
21.9257 | -105.4547
21.9294 | -105.4462
21.9331 | -105.4393
21.9367 | -105.4328
21.9404 | -105.4268
21.9441 | -105.4235
21.9477 | -105.4205
21.9514 | -105.4175
21.9551 | -105.4152
21.9588 | -105.4147
21.9624 | -105.4157
21.9661 | -105.4166
21.9698 | -105.4189
21.9735 | -105.4234
21.9771 | -105.4281
21.9808 | -105.4344
21.9845 | -105.4411
21.9882 | -105.4491
21.9918 | -105.4589
21.9955 | -105.4656
21.9992 | -105.4704
22.0028 | -105.4762
22.0065 | -105.4858
22.0102 | -105.4952
22.0139 | -105.5025
22.0175 | -105.509
22.0212 | -105.5147
22.0249 | -105.5232
22.0286 | -105.5334
22.0322 | -105.5404
22.0359 | -105.5463
22.0396 | -105.5536
22.0433 | -105.5603
22.0469 | -105.5689
22.0506 | -105.5783
22.0543 | -105.5878
22.0579 | -105.5945
22.0616 | -105.5994
22.0653 | -105.6084
22.069 | -105.6161
22.0726 | -105.6222
22.0763 | -105.6261
22.08 | -105.6301
22.0837 | -105.6348
22.0873 | -105.6394
22.091 | -105.6442
22.0947 | -105.6496
22.0984 | -105.6556
22.102 | -105.6624
22.1057 | -105.6686
22.1094 | -105.6726
22.113 | -105.6762
22.1167 | -105.679
22.1204 | -105.6796
22.1241 | -105.6819
22.1277 | -105.6851
22.1314 | -105.6855
22.1351 | -105.6835
22.1388 | -105.6824
22.1424 | -105.6826
22.1461 | -105.6828
22.1498 | -105.6807
22.1535 | -105.681
22.1571 | -105.6803
22.1608 | -105.6806
22.1645 | -105.6793
22.1681 | -105.6786
22.1718 | -105.6782
22.1755 | -105.6774
22.1792 | -105.6735
22.1828 | -105.6686
22.1865 | -105.6665
22.1902 | -105.6641
22.1939 | -105.6599
22.1975 | -105.6579
22.2012 | -105.6587
22.2049 | -105.6573
22.2085 | -105.6538
22.2122 | -105.6501
22.2159 | -105.6465
22.2196 | -105.6446
22.2232 | -105.6421
22.2269 | -105.6394
22.2306 | -105.6383
22.2343 | -105.6359
22.2379 | -105.6342
22.2416 | -105.6322
22.2453 | -105.6303
22.249 | -105.6289
22.2526 | -105.6279
22.2563 | -105.6277
22.26 | -105.6267
22.2636 | -105.6281
22.2673 | -105.6313
22.271 | -105.6339
22.2747 | -105.6356
22.2783 | -105.6353
22.282 | -105.6378
22.2857 | -105.6399
22.2894 | -105.6401
22.293 | -105.6432
22.2967 | -105.6478
22.3004 | -105.6509
22.3041 | -105.6545
22.3077 | -105.6586
22.3114 | -105.6626
22.3151 | -105.67
22.3187 | -105.6772
22.3224 | -105.6834
22.3261 | -105.6891
22.3298 | -105.6945
22.3334 | -105.7021
22.3371 | -105.7117
22.3408 | -105.7203
22.3445 | -105.7268
22.3481 | -105.7339
22.3518 | -105.7398
22.3555 | -105.7452
22.3592 | -105.7536
22.3628 | -105.759
22.3665 | -105.7629
22.3702 | -105.7677
22.3738 | -105.7732
22.3775 | -105.7793
22.3812 | -105.7835
22.3849 | -105.7869
22.3885 | -105.7919
22.3922 | -105.7975
22.3959 | -105.8025
22.3996 | -105.8071
22.4032 | -105.8106
22.4069 | -105.8131
22.4106 | -105.815
22.4143 | -105.8163
22.4179 | -105.816
22.4216 | -105.8164
22.4253 | -105.8176
22.4289 | -105.8127
22.4326 | -105.8041
22.4363 | -105.7897
22.44 | -105.7628
22.4436 | -105.7352
22.4473 | -105.7177
22.451 | -105.7011
22.4547 | -105.685
22.4583 | -105.666
22.462 | -105.6441
22.4657 | -105.6233
22.4694 | -105.6008
22.473 | -105.5783
22.4767 | -105.5522
22.4804 | -105.5238
22.484 | -105.4934
22.4877 | -105.4622
22.4914 | -105.4313
22.4951 | -105.4003
22.4987 | -105.3665
22.5024 | -105.3282
22.5061 | -105.2877
22.5098 | -105.249
22.5134 | -105.2086
22.5171 | -105.1665
22.5208 | -105.1258
22.5245 | -105.0864
22.5281 | -105.0451
22.5318 | -105.004
22.5355 | -104.9652
22.5391 | -104.9278
22.5428 | -104.8891
22.5465 | -104.8512
22.5502 | -104.8152
22.5538 | -104.7812
22.5575 | -104.7497
22.5612 | -104.7186
22.5649 | -104.6887
22.5685 | -104.6614
22.5722 | -104.6399
22.5759 | -104.6217
22.5795 | -104.6036
22.5832 | -104.5895
22.5869 | -104.5786
22.5906 | -104.5708
22.5942 | -104.5664
22.5979 | -104.5665
22.6016 | -104.5693
22.6053 | -104.577
22.6089 | -104.5899
22.6126 | -104.6038
22.6163 | -104.6207
22.62 | -104.6416
22.6236 | -104.6661
22.6273 | -104.6943
22.631 | -104.7254
22.6346 | -104.7571
22.6383 | -104.7924
22.642 | -104.8305
22.6457 | -104.8719
22.6493 | -104.9156
22.653 | -104.9594
22.6567 | -105.0042
22.6604 | -105.0495
22.664 | -105.0969
22.6677 | -105.1471
22.6714 | -105.1981
22.6751 | -105.2479
22.6787 | -105.2966
22.6824 | -105.3448
22.6861 | -105.3932
22.6897 | -105.4416
22.6934 | -105.4885
22.6971 | -105.5326
22.7008 | -105.5768
22.7044 | -105.6202
22.7081 | -105.6605
22.7118 | -105.6998
22.7155 | -105.7368
22.7191 | -105.7726
22.7228 | -105.8074
22.7265 | -105.8397
22.7302 | -105.8688
22.7338 | -105.8966
22.7375 | -105.9235
22.7412 | -105.947
22.7448 | -105.9711
22.7485 | -105.9968
22.7522 | -106.0226
22.7559 | -106.0515
22.7595 | -106.0851
22.7632 | -106.1123
22.7669 | -106.126
22.7706 | -106.1355
22.7742 | -106.1405
22.7779 | -106.1434
22.7816 | -106.1462
22.7853 | -106.1477
22.7889 | -106.1503
22.7926 | -106.1523
22.7963 | -106.1528
22.7999 | -106.1528
22.8036 | -106.154
22.8073 | -106.1548
22.811 | -106.1551
22.8146 | -106.1541
22.8183 | -106.1541
22.822 | -106.1536
22.8257 | -106.152
22.8293 | -106.1539
22.833 | -106.1531
22.8367 | -106.1516
22.8404 | -106.1481
22.844 | -106.1432
22.8477 | -106.1389
22.8514 | -106.1358
22.855 | -106.1339
22.8587 | -106.1299
22.8624 | -106.1263
22.8661 | -106.1245
22.8697 | -106.121
22.8734 | -106.1182
22.8771 | -106.1149
22.8808 | -106.1116
22.8844 | -106.1092
22.8881 | -106.1088
22.8918 | -106.1083
22.8955 | -106.1065
22.8991 | -106.1034
22.9028 | -106.1023
22.9065 | -106.1007
22.9101 | -106.0984
22.9138 | -106.0967
22.9175 | -106.0972
22.9212 | -106.0976
22.9248 | -106.0973
22.9285 | -106.0947
22.9322 | -106.0908
22.9359 | -106.0909
22.9395 | -106.0919
22.9432 | -106.0932
22.9469 | -106.0925
22.9505 | -106.0926
22.9542 | -106.0937
22.9579 | -106.0948
22.9616 | -106.0953
22.9652 | -106.0948
22.9689 | -106.0924
22.9726 | -106.0917
22.9763 | -106.0917
22.9799 | -106.0925
22.9836 | -106.0944
22.9873 | -106.0942
22.991 | -106.0911
22.9946 | -106.0893
22.9983 | -106.0886
23.002 | -106.0877
23.0056 | -106.0882
23.0093 | -106.0872
23.013 | -106.0842
23.0167 | -106.083
23.0203 | -106.0813
23.024 | -106.0788
23.0277 | -106.0751
23.0314 | -106.0708
23.035 | -106.067
23.0387 | -106.0625
23.0424 | -106.0574
23.0461 | -106.051
23.0497 | -106.0431
23.0534 | -106.0375
23.0571 | -106.0314
23.0607 | -106.0249
23.0644 | -106.0197
23.0681 | -106.0111
23.0718 | -106.0033
23.0754 | -105.9957
23.0791 | -105.9868
23.0828 | -105.9792
23.0865 | -105.9717
23.0901 | -105.9638
23.0938 | -105.9552
23.0975 | -105.9467
23.1012 | -105.9389
23.1048 | -105.9325
23.1085 | -105.9241
23.1122 | -105.9179
23.1158 | -105.9129
23.1195 | -105.9081
23.1232 | -105.903
23.1269 | -105.8986
23.1305 | -105.8946
23.1342 | -105.8927
23.1379 | -105.8907
23.1416 | -105.8878
23.1452 | -105.8843
23.1489 | -105.8844
23.1526 | -105.8852
23.1563 | -105.8851
23.1599 | -105.8863
23.1636 | -105.8884
23.1673 | -105.8908
23.1709 | -105.8938
23.1746 | -105.8971
23.1783 | -105.9001
23.182 | -105.9028
23.1856 | -105.9057
23.1893 | -105.908
23.193 | -105.9121
23.1967 | -105.9182
23.2003 | -105.9238
23.204 | -105.9291
23.2077 | -105.9335
23.2113 | -105.9405
23.215 | -105.9447
23.2187 | -105.9475
23.2224 | -105.9507
23.226 | -105.9555
23.2297 | -105.9589
23.2334 | -105.9637
23.2371 | -105.9691
23.2407 | -105.974
23.2444 | -105.977
23.2481 | -105.9806
23.2518 | -105.9848
23.2554 | -105.9884
23.2591 | -105.9917
23.2628 | -105.9939
23.2664 | -105.9975
23.2701 | -106.0021
23.2738 | -106.0064
23.2775 | -106.009
23.2811 | -106.0123
23.2848 | -106.0159
23.2885 | -106.017
23.2922 | -106.0188
23.2958 | -106.0224
23.2995 | -106.0236
23.3032 | -106.0235
23.3069 | -106.0225
23.3105 | -106.0216
23.3142 | -106.0216
23.3179 | -106.0243
23.3215 | -106.0266
23.3252 | -106.0287
23.3289 | -106.0305
23.3326 | -106.0321
23.3362 | -106.0325
23.3399 | -106.034
23.3436 | -106.0363
23.3473 | -106.0362
23.3509 | -106.0362
23.3546 | -106.0358
23.3583 | -106.0365
23.362 | -106.0367
23.3656 | -106.0376
23.3693 | -106.0403
23.373 | -106.0418
23.3766 | -106.0439
23.3803 | -106.0466
23.384 | -106.0484
23.3877 | -106.0506
23.3913 | -106.0527
23.395 | -106.0514
23.3987 | -106.0507
23.4024 | -106.051
23.406 | -106.0503
23.4097 | -106.0504
23.4134 | -106.0498
23.4171 | -106.0494
23.4207 | -106.0485
23.4244 | -106.0491
23.4281 | -106.049
23.4317 | -106.048
23.4354 | -106.0444
23.4391 | -106.0423
23.4428 | -106.0407
23.4464 | -106.0379
23.4501 | -106.0355
23.4538 | -106.0331
23.4575 | -106.0308
23.4611 | -106.0273
23.4648 | -106.0236
23.4685 | -106.0214
23.4722 | -106.0193
23.4758 | -106.0159
23.4795 | -106.0098
23.4832 | -106.0057
23.4868 | -106.0008
23.4905 | -105.9953
23.4942 | -105.9883
23.4979 | -105.9837
23.5015 | -105.979
23.5052 | -105.9728
23.5089 | -105.9681
23.5126 | -105.9634
23.5162 | -105.959
23.5199 | -105.9538
23.5236 | -105.9475
23.5273 | -105.9403
23.5309 | -105.9334
23.5346 | -105.9267
23.5383 | -105.9203
23.5419 | -105.9123
23.5456 | -105.9068
23.5493 | -105.9042
23.553 | -105.9003
23.5566 | -105.8964
23.5603 | -105.8924
23.564 | -105.8866
23.5677 | -105.8811
23.5713 | -105.8755
23.575 | -105.8704
23.5787 | -105.8674
23.5824 | -105.865
23.586 | -105.864
23.5897 | -105.8647
23.5934 | -105.8638
23.597 | -105.8617
23.6007 | -105.8581
23.6044 | -105.8549
23.6081 | -105.854
23.6117 | -105.8531
23.6154 | -105.8533
23.6191 | -105.8549
23.6228 | -105.8558
23.6264 | -105.8575
23.6301 | -105.8581
23.6338 | -105.8585
23.6374 | -105.8601
23.6411 | -105.8602
23.6448 | -105.8602
23.6485 | -105.8593
23.6521 | -105.858
23.6558 | -105.8565
23.6595 | -105.8554
23.6632 | -105.8568
23.6668 | -105.858
23.6705 | -105.859
23.6742 | -105.8592
23.6779 | -105.8599
23.6815 | -105.8608
23.6852 | -105.861
23.6889 | -105.863
23.6925 | -105.8649
23.6962 | -105.8666
23.6999 | -105.8675
23.7036 | -105.8688
23.7072 | -105.8702
23.7109 | -105.8711
23.7146 | -105.8721
23.7183 | -105.8722
23.7219 | -105.8727
23.7256 | -105.8736
23.7293 | -105.8758
23.733 | -105.8793
23.7366 | -105.8807
23.7403 | -105.8825
23.744 | -105.882
23.7476 | -105.881
23.7513 | -105.8821
23.755 | -105.8843
23.7587 | -105.8839
23.7623 | -105.8846
23.766 | -105.8848
23.7697 | -105.8875
23.7734 | -105.8899
23.777 | -105.8916
23.7807 | -105.8935
23.7844 | -105.8962
23.7881 | -105.8989
23.7917 | -105.9025
23.7954 | -105.9073
23.7991 | -105.9092
23.8027 | -105.911
23.8064 | -105.9142
23.8101 | -105.9192
23.8138 | -105.9239
23.8174 | -105.9244
23.8211 | -105.9278
23.8248 | -105.9328
23.8285 | -105.9396
23.8321 | -105.9456
23.8358 | -105.9498
23.8395 | -105.9538
23.8432 | -105.9572
23.8468 | -105.9624
23.8505 | -105.9676
23.8542 | -105.9729
23.8578 | -105.9782
23.8615 | -105.984
23.8652 | -105.99
23.8689 | -105.9968
23.8725 | -106.0038
23.8762 | -106.0094
23.8799 | -106.0143
23.8836 | -106.0185
23.8872 | -106.0219
23.8909 | -106.0243
23.8946 | -106.0275
23.8983 | -106.0304
23.9019 | -106.0333
23.9056 | -106.0377
23.9093 | -106.0426
23.9129 | -106.0463
23.9166 | -106.0469
23.9203 | -106.0462
23.924 | -106.0445
23.9276 | -106.0435
23.9313 | -106.0436
23.935 | -106.045
23.9387 | -106.0451
23.9423 | -106.0455
23.946 | -106.0447
23.9497 | -106.0444
23.9534 | -106.0434
23.957 | -106.0418
23.9607 | -106.0393
23.9644 | -106.034
23.968 | -106.0316
23.9717 | -106.031
23.9754 | -106.0276
23.9791 | -106.0238
23.9827 | -106.0193
23.9864 | -106.0143
23.9901 | -106.0102
23.9938 | -106.0055
23.9974 | -105.9997
24.0011 | -105.9917
24.0048 | -105.9822
24.0084 | -105.9747
24.0121 | -105.9683
24.0158 | -105.959
24.0195 | -105.951
24.0231 | -105.9448
24.0268 | -105.9336
24.0305 | -105.9198
24.0342 | -105.9096
24.0378 | -105.9001
24.0415 | -105.8877
24.0452 | -105.8755
24.0489 | -105.8625
24.0525 | -105.8511
24.0562 | -105.84
24.0599 | -105.8269
24.0635 | -105.8132
24.0672 | -105.8003
24.0709 | -105.7867
24.0746 | -105.7718
24.0782 | -105.7554
24.0819 | -105.7401
24.0856 | -105.7263
24.0893 | -105.7097
24.0929 | -105.6906
24.0966 | -105.6748
24.1003 | -105.6588
24.104 | -105.6447
24.1076 | -105.6272
24.1113 | -105.6091
24.115 | -105.5927
24.1186 | -105.5774
24.1223 | -105.561
24.126 | -105.5465
24.1297 | -105.5312
24.1333 | -105.5158
24.137 | -105.5043
24.1407 | -105.4929
24.1444 | -105.4789
24.148 | -105.4635
24.1517 | -105.4508
24.1554 | -105.441
24.1591 | -105.4296
24.1627 | -105.4183
24.1664 | -105.4074
24.1701 | -105.3967
24.1737 | -105.3878
24.1774 | -105.3801
24.1811 | -105.3745
24.1848 | -105.3687
24.1884 | -105.3629
24.1921 | -105.3571
24.1958 | -105.3522
24.1995 | -105.3492
24.2031 | -105.3459
24.2068 | -105.3415
24.2105 | -105.3383
24.2142 | -105.3351
24.2178 | -105.3344
24.2215 | -105.3359
24.2252 | -105.3368
24.2288 | -105.3378
24.2325 | -105.3384
24.2362 | -105.3387
24.2399 | -105.341
24.2435 | -105.343
24.2472 | -105.3432
24.2509 | -105.3446
24.2546 | -105.3469
24.2582 | -105.3501
24.2619 | -105.3534
24.2656 | -105.3567
24.2692 | -105.3569
24.2729 | -105.3562
24.2766 | -105.3505
24.2803 | -105.3432
24.2839 | -105.3319
24.2876 | -105.3188
24.2913 | -105.3075
24.295 | -105.2995
24.2986 | -105.2933
24.3023 | -105.2855
24.306 | -105.2767
24.3097 | -105.266
24.3133 | -105.2535
24.317 | -105.2412
24.3207 | -105.2287
24.3243 | -105.2151
24.328 | -105.1991
24.3317 | -105.1798
24.3354 | -105.1594
24.339 | -105.1366
24.3427 | -105.115
24.3464 | -105.0943
24.3501 | -105.0713
24.3537 | -105.0458
24.3574 | -105.0207
24.3611 | -104.9935
24.3648 | -104.9654
24.3684 | -104.9404
24.3721 | -104.915
24.3758 | -104.8853
24.3794 | -104.8578
24.3831 | -104.829
24.3868 | -104.8002
24.3905 | -104.773
24.3941 | -104.7465
24.3978 | -104.7192
24.4015 | -104.6934
24.4052 | -104.6698
24.4088 | -104.6469
24.4125 | -104.6236
24.4162 | -104.601
24.4199 | -104.5816
24.4235 | -104.5637
24.4272 | -104.5473
24.4309 | -104.5312
24.4345 | -104.5168
24.4382 | -104.5038
24.4419 | -104.492
24.4456 | -104.4811
24.4492 | -104.4709
24.4529 | -104.4637
24.4566 | -104.4587
24.4603 | -104.4537
24.4639 | -104.4507
24.4676 | -104.4501
24.4713 | -104.4507
24.475 | -104.4534
24.4786 | -104.4567
24.4823 | -104.4603
24.486 | -104.4656
24.4896 | -104.4743
24.4933 | -104.4853
24.497 | -104.4965
24.5007 | -104.5079
24.5043 | -104.5195
24.508 | -104.5299
24.5117 | -104.5403
24.5154 | -104.5519
24.519 | -104.5655
24.5227 | -104.5793
24.5264 | -104.5918
24.5301 | -104.605
24.5337 | -104.6191
24.5374 | -104.6326
24.5411 | -104.6453
24.5447 | -104.6583
24.5484 | -104.6726
24.5521 | -104.6865
24.5558 | -104.6988
24.5594 | -104.7092
24.5631 | -104.7186
24.5668 | -104.7269
24.5705 | -104.732
24.5741 | -104.7269
24.5778 | -104.7123
24.5815 | -104.6644
24.5852 | -104.5472
24.5888 | -104.4425
24.5925 | -104.3754
24.5962 | -104.3192
24.5998 | -104.269
24.6035 | -104.2171
24.6072 | -104.1607
24.6109 | -104.1008
24.6145 | -104.0356
24.6182 | -103.9624
24.6219 | -103.8876
24.6256 | -103.8109
24.6292 | -103.7316
24.6329 | -103.6469
24.6366 | -103.5566
24.6403 | -103.4606
24.6439 | -103.3604
24.6476 | -103.2597
24.6513 | -103.1582
24.6549 | -103.0541
24.6586 | -102.949
24.6623 | -102.8433
24.666 | -102.7365
24.6696 | -102.6292
24.6733 | -102.5232
24.677 | -102.419
24.6807 | -102.316
24.6843 | -102.2144
24.688 | -102.1144
24.6917 | -102.0166
24.6953 | -101.9221
24.699 | -101.8321
24.7027 | -101.7449
24.7064 | -101.6622
24.71 | -101.5836
24.7137 | -101.5091
24.7174 | -101.4394
24.7211 | -101.375
24.7247 | -101.3159
24.7284 | -101.2617
24.7321 | -101.2134
24.7358 | -101.1716
24.7394 | -101.1351
24.7431 | -101.1045
24.7468 | -101.08
24.7504 | -101.0615
24.7541 | -101.0488
24.7578 | -101.0425
24.7615 | -101.0429
24.7651 | -101.0491
24.7688 | -101.0618
24.7725 | -101.0799
24.7762 | -101.1041
24.7798 | -101.1341
24.7835 | -101.17
24.7872 | -101.2128
24.7909 | -101.2602
24.7945 | -101.3138
24.7982 | -101.3738
24.8019 | -101.4385
24.8055 | -101.5076
24.8092 | -101.5817
24.8129 | -101.6602
24.8166 | -101.7429
24.8202 | -101.8296
24.8239 | -101.9198
24.8276 | -102.014
24.8313 | -102.1115
24.8349 | -102.2105
24.8386 | -102.3111
24.8423 | -102.4137
24.846 | -102.5184
24.8496 | -102.624
24.8533 | -102.7292
24.857 | -102.8337
24.8606 | -102.94
24.8643 | -103.0459
24.868 | -103.1496
24.8717 | -103.2517
24.8753 | -103.3518
24.879 | -103.4487
24.8827 | -103.5437
24.8864 | -103.6371
24.89 | -103.7274
24.8937 | -103.8131
24.8974 | -103.8937
24.9011 | -103.9722
24.9047 | -104.0003
24.9084 | -104.0696
24.9121 | -104.1372
24.9157 | -104.2015
24.9194 | -104.2614
24.9231 | -104.3192
24.9268 | -104.3796
24.9304 | -104.4415
24.9341 | -104.5233
24.9378 | -104.6532
24.9415 | -104.7484
24.9451 | -104.787
24.9488 | -104.8022
24.9525 | -104.8044
24.9562 | -104.8037
24.9598 | -104.8014
24.9635 | -104.7984
24.9672 | -104.7957
24.9708 | -104.7965
24.9745 | -104.7956
24.9782 | -104.7908
24.9819 | -104.7849
24.9855 | -104.7797
24.9892 | -104.7731
24.9929 | -104.7648
24.9966 | -104.7565
25.0002 | -104.7473
25.0039 | -104.7379
25.0076 | -104.7289
25.0112 | -104.7187
25.0149 | -104.7093
25.0186 | -104.7017
25.0223 | -104.6953
25.0259 | -104.6882
25.0296 | -104.6792
25.0333 | -104.6711
25.037 | -104.6644
25.0406 | -104.657
25.0443 | -104.6512
25.048 | -104.6429
25.0517 | -104.6337
25.0553 | -104.6259
25.059 | -104.6177
25.0627 | -104.6105
25.0663 | -104.6026
25.07 | -104.5949
25.0737 | -104.589
25.0774 | -104.5833
25.081 | -104.5786
25.0847 | -104.5729
25.0884 | -104.5668
25.0921 | -104.5622
25.0957 | -104.5566
25.0994 | -104.5502
25.1031 | -104.5431
25.1068 | -104.5379
25.1104 | -104.5331
25.1141 | -104.5285
25.1178 | -104.5254
25.1214 | -104.5218
25.1251 | -104.5203
25.1288 | -104.5187
25.1325 | -104.5165
25.1361 | -104.515
25.1398 | -104.5131
25.1435 | -104.5108
25.1472 | -104.5083
25.1508 | -104.5048
25.1545 | -104.5025
25.1582 | -104.5001
25.1619 | -104.4962
25.1655 | -104.492
25.1692 | -104.4904
25.1729 | -104.4892
25.1765 | -104.4882
25.1802 | -104.4853
25.1839 | -104.4838
25.1876 | -104.484
25.1912 | -104.4855
25.1949 | -104.4862
25.1986 | -104.4857
25.2023 | -104.4863
25.2059 | -104.4883
25.2096 | -104.4921
25.2133 | -104.4913
25.217 | -104.4904
25.2206 | -104.4914
25.2243 | -104.4922
25.228 | -104.4924
25.2316 | -104.4928
25.2353 | -104.4935
25.239 | -104.4942
25.2427 | -104.4934
25.2463 | -104.4934
25.25 | -104.4952
25.2537 | -104.4973
25.2574 | -104.4989
25.261 | -104.5001
25.2647 | -104.5007
25.2684 | -104.4996
25.2721 | -104.4988
25.2757 | -104.4982
25.2794 | -104.4975
25.2831 | -104.4977
25.2867 | -104.4978
25.2904 | -104.4979
25.2941 | -104.4989
25.2978 | -104.4999
25.3014 | -104.5004
25.3051 | -104.5012
25.3088 | -104.501
25.3125 | -104.5013
25.3161 | -104.5019
25.3198 | -104.5023
25.3235 | -104.5041
25.3271 | -104.5043
25.3308 | -104.5045
25.3345 | -104.5063
25.3382 | -104.5068
25.3418 | -104.5059
25.3455 | -104.5042
25.3492 | -104.503
25.3529 | -104.5031
25.3565 | -104.5031
25.3602 | -104.5017
25.3639 | -104.5011
25.3676 | -104.5011
25.3712 | -104.4987
25.3749 | -104.4952
25.3786 | -104.4922
25.3822 | -104.4905
25.3859 | -104.4897
25.3896 | -104.4883
25.3933 | -104.4858
25.3969 | -104.4837
25.4006 | -104.4807
25.4043 | -104.4799
25.408 | -104.4781
25.4116 | -104.4773
25.4153 | -104.4774
25.419 | -104.4759
25.4227 | -104.4757
25.4263 | -104.4763
25.43 | -104.4737
25.4337 | -104.4752
25.4373 | -104.4735
25.441 | -104.4727
25.4447 | -104.4727
25.4484 | -104.4734
25.452 | -104.4729
25.4557 | -104.4743
25.4594 | -104.4779
25.4631 | -104.4821
25.4667 | -104.4849
25.4704 | -104.4858
25.4741 | -104.487
25.4778 | -104.4887
25.4814 | -104.4905
25.4851 | -104.4931
25.4888 | -104.4955
25.4924 | -104.499
25.4961 | -104.501
25.4998 | -104.5037
25.5035 | -104.509
25.5071 | -104.5136
25.5108 | -104.5179
25.5145 | -104.5222
25.5182 | -104.5276
25.5218 | -104.5341
25.5255 | -104.5403
25.5292 | -104.5454
25.5329 | -104.5486
25.5365 | -104.552
25.5402 | -104.5553
25.5439 | -104.5576
25.5475 | -104.5613
25.5512 | -104.5643
25.5549 | -104.5674
25.5586 | -104.5697
25.5622 | -104.5719
25.5659 | -104.5752
25.5696 | -104.5789
25.5733 | -104.5832
25.5769 | -104.5858
25.5806 | -104.5869
25.5843 | -104.5892
25.588 | -104.5934
25.5916 | -104.5963
25.5953 | -104.5985
25.599 | -104.6008
25.6026 | -104.6039
25.6063 | -104.6068
25.61 | -104.6102
25.6137 | -104.6115
25.6173 | -104.6139
25.621 | -104.618
25.6247 | -104.6218
25.6284 | -104.6258
25.632 | -104.6275
25.6357 | -104.6296
25.6394 | -104.6311
25.6431 | -104.6324
25.6467 | -104.6334
25.6504 | -104.6348
25.6541 | -104.6344
25.6577 | -104.6334
25.6614 | -104.6336
25.6651 | -104.6342
25.6688 | -104.635
25.6724 | -104.636
25.6761 | -104.6371
25.6798 | -104.6374
25.6835 | -104.6373
25.6871 | -104.6385
25.6908 | -104.6392
25.6945 | -104.6381
25.6981 | -104.6386
25.7018 | -104.6391
25.7055 | -104.6393
25.7092 | -104.6403
25.7128 | -104.6394
25.7165 | -104.6389
25.7202 | -104.6379
25.7239 | -104.637
25.7275 | -104.6371
25.7312 | -104.6397
25.7349 | -104.6416
25.7386 | -104.6421
25.7422 | -104.6412
25.7459 | -104.6419
25.7496 | -104.6425
25.7532 | -104.6436
25.7569 | -104.6442
25.7606 | -104.6444
25.7643 | -104.6452
25.7679 | -104.6463
25.7716 | -104.647
25.7753 | -104.6457
25.779 | -104.6464
25.7826 | -104.6483
25.7863 | -104.6514
25.79 | -104.6538
25.7937 | -104.6564
25.7973 | -104.6589
25.801 | -104.6635
25.8047 | -104.6668
25.8083 | -104.6698
25.812 | -104.6753
25.8157 | -104.6816
25.8194 | -104.6861
25.823 | -104.6898
25.8267 | -104.6947
25.8304 | -104.7006
25.8341 | -104.708
25.8377 | -104.715
25.8414 | -104.7223
25.8451 | -104.731
25.8488 | -104.7387
25.8524 | -104.7469
25.8561 | -104.7543
25.8598 | -104.763
25.8634 | -104.7697
25.8671 | -104.7759
25.8708 | -104.7841
25.8745 | -104.7936
25.8781 | -104.8031
25.8818 | -104.8117
25.8855 | -104.8203
25.8892 | -104.8283
25.8928 | -104.8353
25.8965 | -104.8447
25.9002 | -104.8527
25.9039 | -104.8619
25.9075 | -104.8703
25.9112 | -104.8762
25.9149 | -104.8823
25.9185 | -104.8873
25.9222 | -104.8935
25.9259 | -104.8996
25.9296 | -104.9045
25.9332 | -104.9086
25.9369 | -104.9125
25.9406 | -104.9166
25.9443 | -104.9196
25.9479 | -104.9213
25.9516 | -104.9223
25.9553 | -104.9241
25.959 | -104.9245
25.9626 | -104.9239
25.9663 | -104.9202
25.97 | -104.9204
25.9736 | -104.9182
25.9773 | -104.9147
25.981 | -104.9134
25.9847 | -104.9112
25.9883 | -104.9075
25.992 | -104.9043
25.9957 | -104.9014
25.9994 | -104.8991
26.003 | -104.8949
26.0067 | -104.892
26.0104 | -104.8879
26.0141 | -104.8833
26.0177 | -104.8776
26.0214 | -104.8726
26.0251 | -104.8688
26.0287 | -104.8655
26.0324 | -104.8593
26.0361 | -104.8513
26.0398 | -104.8427
26.0434 | -104.8349
26.0471 | -104.8269
26.0508 | -104.8183
26.0545 | -104.8104
26.0581 | -104.8033
26.0618 | -104.7972
26.0655 | -104.7894
26.0691 | -104.7809
26.0728 | -104.7732
26.0765 | -104.7659
26.0802 | -104.7587
26.0838 | -104.7517
26.0875 | -104.746
26.0912 | -104.7397
26.0949 | -104.7325
26.0985 | -104.7251
26.1022 | -104.7193
26.1059 | -104.7131
26.1096 | -104.708
26.1132 | -104.702
26.1169 | -104.6966
26.1206 | -104.6912
26.1242 | -104.6864
26.1279 | -104.6821
26.1316 | -104.6789
26.1353 | -104.6748
26.1389 | -104.6691
26.1426 | -104.6659
26.1463 | -104.6629
26.15 | -104.6599
26.1536 | -104.6557
26.1573 | -104.6522
26.161 | -104.6488
26.1647 | -104.6465
26.1683 | -104.6448
26.172 | -104.6444
26.1757 | -104.645
26.1793 | -104.6448
26.183 | -104.6443
26.1867 | -104.6458
26.1904 | -104.6482
26.194 | -104.6504
26.1977 | -104.6522
26.2014 | -104.655
26.2051 | -104.6587
26.2087 | -104.6613
26.2124 | -104.6648
26.2161 | -104.6697
26.2198 | -104.6745
26.2234 | -104.6813
26.2271 | -104.6885
26.2308 | -104.6965
26.2344 | -104.7037
26.2381 | -104.7076
26.2418 | -104.7089
26.2455 | -104.7097
26.2491 | -104.7113
26.2528 | -104.7161
26.2565 | -104.7227
26.2602 | -104.7293
26.2638 | -104.7356
26.2675 | -104.7425
26.2712 | -104.7501
26.2749 | -104.7566
26.2785 | -104.7629
26.2822 | -104.7679
26.2859 | -104.773
26.2895 | -104.7761
26.2932 | -104.7789
26.2969 | -104.7828
26.3006 | -104.7867
26.3042 | -104.7913
26.3079 | -104.7957
26.3116 | -104.7991
26.3153 | -104.8027
26.3189 | -104.8055
26.3226 | -104.8072
26.3263 | -104.8085
26.33 | -104.8097
26.3336 | -104.8104
26.3373 | -104.8111
26.341 | -104.8117
26.3446 | -104.8123
26.3483 | -104.8122
26.352 | -104.8117
26.3557 | -104.8121
26.3593 | -104.8122
26.363 | -104.8118
26.3667 | -104.8109
26.3704 | -104.8079
26.374 | -104.8061
26.3777 | -104.8053
26.3814 | -104.8054
26.385 | -104.8056
26.3887 | -104.8046
26.3924 | -104.8036
26.3961 | -104.8027
26.3997 | -104.8014
26.4034 | -104.7995
26.4071 | -104.7998
26.4108 | -104.7993
26.4144 | -104.7978
26.4181 | -104.7981
26.4218 | -104.7991
26.4255 | -104.7985
26.4291 | -104.7976
26.4328 | -104.799
26.4365 | -104.8012
26.4401 | -104.8051
26.4438 | -104.8075
26.4475 | -104.8086
26.4512 | -104.8114
26.4548 | -104.8141
26.4585 | -104.8161
26.4622 | -104.8195
26.4659 | -104.8242
26.4695 | -104.8281
26.4732 | -104.8321
26.4769 | -104.8375
26.4806 | -104.8436
26.4842 | -104.8483
26.4879 | -104.8529
26.4916 | -104.8583
26.4952 | -104.8647
26.4989 | -104.8705
26.5026 | -104.8753
26.5063 | -104.8792
26.5099 | -104.883
26.5136 | -104.8895
26.5173 | -104.8957
26.521 | -104.9
26.5246 | -104.9036
26.5283 | -104.909
26.532 | -104.9144
26.5357 | -104.9211
26.5393 | -104.9256
26.543 | -104.9284
26.5467 | -104.9312
26.5503 | -104.9353
26.554 | -104.9394
26.5577 | -104.9429
26.5614 | -104.9461
26.565 | -104.947
26.5687 | -104.9487
26.5724 | -104.9525
26.5761 | -104.9566
26.5797 | -104.9599
26.5834 | -104.9631
26.5871 | -104.964
26.5908 | -104.9651
26.5944 | -104.9683
26.5981 | -104.972
26.6018 | -104.9745
26.6054 | -104.9756
26.6091 | -104.979
26.6128 | -104.9832
26.6165 | -104.9889
26.6201 | -104.9966
26.6238 | -105.0033
26.6275 | -105.0053
26.6312 | -105.0061
26.6348 | -105.0058
26.6385 | -105.0059
26.6422 | -105.0077
26.6459 | -105.0085
26.6495 | -105.0114
26.6532 | -105.0149
26.6569 | -105.0188
26.6605 | -105.0234
26.6642 | -105.0292
26.6679 | -105.034
26.6716 | -105.0376
26.6752 | -105.0416
26.6789 | -105.0444
26.6826 | -105.0481
26.6863 | -105.0533
26.6899 | -105.058
26.6936 | -105.0626
26.6973 | -105.0678
26.7009 | -105.0732
26.7046 | -105.079
26.7083 | -105.0847
26.712 | -105.0903
26.7156 | -105.0964
26.7193 | -105.1032
26.723 | -105.1082
26.7267 | -105.114
26.7303 | -105.1214
26.734 | -105.1291
26.7377 | -105.1366
26.7414 | -105.1434
26.745 | -105.1532
26.7487 | -105.1614
26.7524 | -105.168
26.756 | -105.1757
26.7597 | -105.1823
26.7634 | -105.1893
26.7671 | -105.1963
26.7707 | -105.2042
26.7744 | -105.2123
26.7781 | -105.2229
26.7818 | -105.2319
26.7854 | -105.2413
26.7891 | -105.2531
26.7928 | -105.2623
26.7965 | -105.2704
26.8001 | -105.2796
26.8038 | -105.2882
26.8075 | -105.295
26.8111 | -105.3018
26.8148 | -105.3072
26.8185 | -105.3137
26.8222 | -105.3209
26.8258 | -105.3266
26.8295 | -105.3297
26.8332 | -105.3338
26.8369 | -105.3375
26.8405 | -105.3403
26.8442 | -105.3449
26.8479 | -105.3494
26.8516 | -105.3532
26.8552 | -105.3556
26.8589 | -105.3592
26.8626 | -105.3626
26.8662 | -105.3675
26.8699 | -105.3714
26.8736 | -105.3726
26.8773 | -105.3757
26.8809 | -105.3797
26.8846 | -105.3831
26.8883 | -105.3866
26.892 | -105.3887
26.8956 | -105.3895
26.8993 | -105.3912
26.903 | -105.3903
26.9067 | -105.3904
26.9103 | -105.3923
26.914 | -105.3919
26.9177 | -105.3921
26.9213 | -105.3938
26.925 | -105.396
26.9287 | -105.3967
26.9324 | -105.3963
26.936 | -105.3978
26.9397 | -105.4011
26.9434 | -105.4058
26.9471 | -105.4093
26.9507 | -105.4125
26.9544 | -105.4156
26.9581 | -105.4199
26.9618 | -105.4234
26.9654 | -105.4282
26.9691 | -105.4328
26.9728 | -105.4373
26.9764 | -105.4408
26.9801 | -105.4435
26.9838 | -105.4491
26.9875 | -105.4554
26.9911 | -105.4631
26.9948 | -105.4698
26.9985 | -105.4752
27.0022 | -105.4819
27.0058 | -105.489
27.0095 | -105.4953
27.0132 | -105.5027
27.0169 | -105.5101
27.0205 | -105.5157
27.0242 | -105.5222
27.0279 | -105.5283
27.0315 | -105.534
27.0352 | -105.5389
27.0389 | -105.5441
27.0426 | -105.5496
27.0462 | -105.5544
27.0499 | -105.5602
27.0536 | -105.5654
27.0573 | -105.5703
27.0609 | -105.5769
27.0646 | -105.5823
27.0683 | -105.5862
27.072 | -105.5912
27.0756 | -105.5954
27.0793 | -105.6016
27.083 | -105.6066
27.0866 | -105.609
27.0903 | -105.6125
27.094 | -105.6154
27.0977 | -105.6175
27.1013 | -105.6209
27.105 | -105.6256
27.1087 | -105.629
27.1124 | -105.63
27.116 | -105.6304
27.1197 | -105.6325
27.1234 | -105.6359
27.127 | -105.6381
27.1307 | -105.6404
27.1344 | -105.6423
27.1381 | -105.6433
27.1417 | -105.6446
27.1454 | -105.6471
27.1491 | -105.6499
27.1528 | -105.6497
27.1564 | -105.6486
27.1601 | -105.6475
27.1638 | -105.6468
27.1675 | -105.6458
27.1711 | -105.6435
27.1748 | -105.6422
27.1785 | -105.6415
27.1821 | -105.6405
27.1858 | -105.6402
27.1895 | -105.6391
27.1932 | -105.6392
27.1968 | -105.6391
27.2005 | -105.6391
27.2042 | -105.64
27.2079 | -105.6401
27.2115 | -105.6422
27.2152 | -105.6443
27.2189 | -105.6446
27.2226 | -105.6443
27.2262 | -105.6427
27.2299 | -105.6407
27.2336 | -105.6383
27.2372 | -105.6373
27.2409 | -105.6359
27.2446 | -105.6341
27.2483 | -105.6295
27.2519 | -105.6272
27.2556 | -105.6237
27.2593 | -105.6197
27.263 | -105.6154
27.2666 | -105.611
27.2703 | -105.6076
27.274 | -105.6047
27.2777 | -105.6018
27.2813 | -105.5988
27.285 | -105.598
27.2887 | -105.5958
27.2923 | -105.5924
27.296 | -105.5888
27.2997 | -105.5858
27.3034 | -105.5836
27.307 | -105.5792
27.3107 | -105.5753
27.3144 | -105.5723
27.3181 | -105.568
27.3217 | -105.5635
27.3254 | -105.5604
27.3291 | -105.5578
27.3328 | -105.5547
27.3364 | -105.5519
27.3401 | -105.5491
27.3438 | -105.5475
27.3474 | -105.5464
27.3511 | -105.5434
27.3548 | -105.5409
27.3585 | -105.5395
27.3621 | -105.5397
27.3658 | -105.5388
27.3695 | -105.5362
27.3732 | -105.5352
27.3768 | -105.5356
27.3805 | -105.537
27.3842 | -105.5373
27.3879 | -105.5391
27.3915 | -105.5412
27.3952 | -105.543
27.3989 | -105.5442
27.4025 | -105.5462
27.4062 | -105.5471
27.4099 | -105.5481
27.4136 | -105.5497
27.4172 | -105.5509
27.4209 | -105.5539
27.4246 | -105.5584
27.4283 | -105.5623
27.4319 | -105.5657
27.4356 | -105.5691
27.4393 | -105.5738
27.4429 | -105.5782
27.4466 | -105.5813
27.4503 | -105.5856
27.454 | -105.5905
27.4576 | -105.5935
27.4613 | -105.5977
27.465 | -105.6012
27.4687 | -105.6061
27.4723 | -105.612
27.476 | -105.6162
27.4797 | -105.6208
27.4834 | -105.6251
27.487 | -105.6282
27.4907 | -105.6322
27.4944 | -105.6353
27.498 | -105.6374
27.5017 | -105.6391
27.5054 | -105.6414
27.5091 | -105.6427
27.5127 | -105.6442
27.5164 | -105.6467
27.5201 | -105.6488
27.5238 | -105.6503
27.5274 | -105.6533
27.5311 | -105.6543
27.5348 | -105.654
27.5385 | -105.6522
27.5421 | -105.6526
27.5458 | -105.6544
27.5495 | -105.6541
27.5531 | -105.6491
27.5568 | -105.6475
27.5605 | -105.6468
27.5642 | -105.645
27.5678 | -105.6437
27.5715 | -105.6426
27.5752 | -105.6403
27.5789 | -105.6382
27.5825 | -105.6351
27.5862 | -105.6328
27.5899 | -105.6279
27.5936 | -105.6228
27.5972 | -105.6182
27.6009 | -105.6131
27.6046 | -105.6066
27.6082 | -105.6014
27.6119 | -105.5968
27.6156 | -105.5923
27.6193 | -105.5882
27.6229 | -105.5852
27.6266 | -105.5824
27.6303 | -105.5798
27.634 | -105.5758
27.6376 | -105.5708
27.6413 | -105.5642
27.645 | -105.5598
27.6487 | -105.5552
27.6523 | -105.5505
27.656 | -105.5448
27.6597 | -105.5374
27.6633 | -105.5311
27.667 | -105.524
27.6707 | -105.5169
27.6744 | -105.5112
27.678 | -105.5041
27.6817 | -105.4974
27.6854 | -105.4929
27.6891 | -105.4865
27.6927 | -105.4792
27.6964 | -105.4733
27.7001 | -105.4659
27.7038 | -105.4566
27.7074 | -105.4504
27.7111 | -105.4423
27.7148 | -105.4336
27.7184 | -105.4253
27.7221 | -105.4164
27.7258 | -105.4095
27.7295 | -105.4025
27.7331 | -105.3948
27.7368 | -105.3875
27.7405 | -105.3794
27.7442 | -105.3712
27.7478 | -105.3631
27.7515 | -105.3549
27.7552 | -105.3468
27.7588 | -105.3385
27.7625 | -105.3294
27.7662 | -105.3209
27.7699 | -105.3111
27.7735 | -105.3021
27.7772 | -105.2937
27.7809 | -105.2865
27.7846 | -105.2788
27.7882 | -105.2721
27.7919 | -105.2657
27.7956 | -105.2601
27.7993 | -105.2531
27.8029 | -105.2464
27.8066 | -105.2381
27.8103 | -105.2282
27.8139 | -105.2204
27.8176 | -105.2123
27.8213 | -105.2058
27.825 | -105.1993
27.8286 | -105.1933
27.8323 | -105.1872
27.836 | -105.1819
27.8397 | -105.1761
27.8433 | -105.1707
27.847 | -105.1662
27.8507 | -105.1607
27.8544 | -105.1543
27.858 | -105.1454
27.8617 | -105.1394
27.8654 | -105.1348
27.869 | -105.1285
27.8727 | -105.1219
27.8764 | -105.1166
27.8801 | -105.1106
27.8837 | -105.1052
27.8874 | -105.1013
27.8911 | -105.0971
27.8948 | -105.0933
27.8984 | -105.0897
27.9021 | -105.0847
27.9058 | -105.0786
27.9095 | -105.074
27.9131 | -105.0703
27.9168 | -105.0656
27.9205 | -105.0618
27.9241 | -105.0539
27.9278 | -105.0428
27.9315 | -105.0317
27.9352 | -105.0226
27.9388 | -105.0164
27.9425 | -105.0109
27.9462 | -105.0045
27.9499 | -104.9975
27.9535 | -104.9904
27.9572 | -104.9839
27.9609 | -104.9768
27.9646 | -104.9688
27.9682 | -104.9615
27.9719 | -104.955
27.9756 | -104.9482
27.9792 | -104.9399
27.9829 | -104.9334
27.9866 | -104.9258
27.9903 | -104.9188
27.9939 | -104.9123
27.9976 | -104.9057
28.0013 | -104.8972