## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 275 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 280
Hood | 02-92B/P | 218
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1010.1 | hPa |  | PRC PRO 01
Humidity | 60 | %/Hr
Internal temperature of calibration bench | 23.95 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1016 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.3 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.98 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.03 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3066 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 31/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 21/08/2018
Technical Manager | Anthony HUE | Approval |  | 30/08/2018
CEO | Robert GLEMOT | Approval |  | 30/08/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4657 | 57.6626
0.012209 | -36.0218 | 38.5087
0.018313 | -34.9537 | 28.3536
0.024417 | -34.4841 | 22.089
0.030521 | -34.3701 | 18.3175
0.036626 | -34.2239 | 15.254
0.04273 | -34.1272 | 12.8825
0.048834 | -34.0797 | 11.3571
0.054938 | -34.0137 | 10.2166
0.061043 | -33.9523 | 9.1287
0.067147 | -33.9417 | 8.3054
0.073251 | -33.9426 | 7.5448
0.079355 | -33.9609 | 6.8828
0.08546 | -33.9745 | 6.2883
0.091564 | -33.969 | 5.8466
0.097668 | -33.9499 | 5.596
0.10377 | -33.9273 | 5.3482
0.10988 | -33.9095 | 5.0711
0.11598 | -33.8999 | 4.7788
0.12209 | -33.9013 | 4.4952
0.12819 | -33.9109 | 4.2115
0.13429 | -33.9216 | 3.9231
0.1404 | -33.9249 | 3.6442
0.1465 | -33.9269 | 3.3859
0.15261 | -33.9243 | 3.1657
0.15871 | -33.918 | 3.0038
0.16482 | -33.91 | 2.902
0.17092 | -33.9074 | 2.8637
0.17702 | -33.9051 | 2.7921
0.18313 | -33.9038 | 2.7139
0.18923 | -33.903 | 2.6047
0.19534 | -33.9013 | 2.4774
0.20144 | -33.8996 | 2.3347
0.20754 | -33.8964 | 2.214
0.21365 | -33.8951 | 2.0995
0.21975 | -33.8943 | 2.002
0.22586 | -33.8951 | 1.9154
0.23196 | -33.8973 | 1.8439
0.23807 | -33.9011 | 1.7823
0.24417 | -33.905 | 1.7245
0.25027 | -33.9088 | 1.6629
0.25638 | -33.9111 | 1.6017
0.26248 | -33.9128 | 1.5355
0.26859 | -33.9136 | 1.4707
0.27469 | -33.9139 | 1.4001
0.2808 | -33.9148 | 1.3272
0.2869 | -33.9155 | 1.257
0.293 | -33.9163 | 1.1908
0.29911 | -33.9171 | 1.1315
0.30521 | -33.9179 | 1.0731
0.31132 | -33.9185 | 1.0346
0.31742 | -33.9196 | 0.99208
0.32353 | -33.9197 | 0.95727
0.32963 | -33.9205 | 0.92817
0.33573 | -33.9213 | 0.90767
0.34184 | -33.9226 | 0.8861
0.34794 | -33.9245 | 0.86196
0.35405 | -33.9266 | 0.83768
0.36015 | -33.9289 | 0.81207
0.36626 | -33.9308 | 0.78431
0.37236 | -33.9327 | 0.74725
0.37846 | -33.9347 | 0.70255
0.38457 | -33.9355 | 0.64484
0.39067 | -33.9364 | 0.59771
0.39678 | -33.9369 | 0.54535
0.40288 | -33.9373 | 0.49082
0.40899 | -33.9371 | 0.43857
0.41509 | -33.9364 | 0.38741
0.42119 | -33.9363 | 0.34425
0.4273 | -33.9362 | 0.29803
0.4334 | -33.9361 | 0.25208
0.43951 | -33.9364 | 0.20915
0.44561 | -33.9371 | 0.16986
0.45172 | -33.9386 | 0.13882
0.45782 | -33.9393 | 0.10225
0.46392 | -33.9404 | 0.066608
0.47003 | -33.9414 | 0.034702
0.47613 | -33.9421 | 0.0011246
0.48224 | -33.9424 | -0.033172
0.48834 | -33.9413 | -0.067795
0.49445 | -33.9409 | -0.10283
0.50055 | -33.9402 | -0.13467
0.50665 | -33.9393 | -0.16871
0.51276 | -33.9382 | -0.20264
0.51886 | -33.937 | -0.23656
0.52497 | -33.9357 | -0.26089
0.53107 | -33.9348 | -0.29037
0.53717 | -33.934 | -0.31863
0.54328 | -33.9331 | -0.3416
0.54938 | -33.9328 | -0.36406
0.55549 | -33.9328 | -0.38506
0.56159 | -33.9339 | -0.39668
0.5677 | -33.9341 | -0.41134
0.5738 | -33.9351 | -0.42765
0.5799 | -33.9359 | -0.44263
0.58601 | -33.937 | -0.45277
0.59211 | -33.9386 | -0.45909
0.59822 | -33.94 | -0.46868
0.60432 | -33.9414 | -0.48071
0.61043 | -33.9427 | -0.48857
0.61653 | -33.9439 | -0.49884
0.62263 | -33.9454 | -0.50817
0.62874 | -33.946 | -0.51864
0.63484 | -33.9469 | -0.52784
0.64095 | -33.9473 | -0.53793
0.64705 | -33.9477 | -0.54745
0.65316 | -33.9478 | -0.55973
0.65926 | -33.9478 | -0.57537
0.66536 | -33.9467 | -0.59252
0.67147 | -33.9462 | -0.60537
0.67757 | -33.9458 | -0.61944
0.68368 | -33.945 | -0.63808
0.68978 | -33.9441 | -0.65503
0.69589 | -33.9431 | -0.67378
0.70199 | -33.9421 | -0.69634
0.70809 | -33.9411 | -0.71532
0.7142 | -33.9404 | -0.73565
0.7203 | -33.9392 | -0.75471
0.72641 | -33.9385 | -0.77716
0.73251 | -33.9377 | -0.79715
0.73862 | -33.9379 | -0.81904
0.74472 | -33.9375 | -0.83916
0.75082 | -33.9372 | -0.86133
0.75693 | -33.9375 | -0.88331
0.76303 | -33.938 | -0.90546
0.76914 | -33.9398 | -0.92628
0.77524 | -33.9407 | -0.94614
0.78135 | -33.9419 | -0.96646
0.78745 | -33.9433 | -0.98655
0.79355 | -33.9445 | -1.0066
0.79966 | -33.9462 | -1.0263
0.80576 | -33.9484 | -1.0463
0.81187 | -33.95 | -1.0681
0.81797 | -33.9519 | -1.0897
0.82408 | -33.9535 | -1.1104
0.83018 | -33.9552 | -1.1312
0.83628 | -33.957 | -1.1498
0.84239 | -33.9581 | -1.1704
0.84849 | -33.9595 | -1.1922
0.8546 | -33.9608 | -1.2122
0.8607 | -33.962 | -1.233
0.86681 | -33.963 | -1.252
0.87291 | -33.9639 | -1.2708
0.87901 | -33.9643 | -1.2887
0.88512 | -33.9649 | -1.3069
0.89122 | -33.9655 | -1.3251
0.89733 | -33.9659 | -1.3407
0.90343 | -33.9662 | -1.3582
0.90953 | -33.9663 | -1.3723
0.91564 | -33.9662 | -1.3879
0.92174 | -33.9664 | -1.4027
0.92785 | -33.9664 | -1.417
0.93395 | -33.9664 | -1.4306
0.94006 | -33.9665 | -1.4432
0.94616 | -33.9661 | -1.4546
0.95226 | -33.9662 | -1.467
0.95837 | -33.9662 | -1.4805
0.96447 | -33.9663 | -1.493
0.97058 | -33.9661 | -1.506
0.97668 | -33.9658 | -1.5189
0.98279 | -33.9655 | -1.5314
0.98889 | -33.9653 | -1.5417
0.99499 | -33.9652 | -1.5547
1.0011 | -33.9651 | -1.5651
1.0072 | -33.9646 | -1.5764
1.0133 | -33.9643 | -1.5869
1.0194 | -33.9638 | -1.5996
1.0255 | -33.9635 | -1.613
1.0316 | -33.9629 | -1.6247
1.0377 | -33.9624 | -1.6381
1.0438 | -33.9619 | -1.6506
1.0499 | -33.9611 | -1.6658
1.056 | -33.9604 | -1.68
1.0621 | -33.9597 | -1.6942
1.0682 | -33.9593 | -1.7066
1.0743 | -33.9586 | -1.7219
1.0805 | -33.9581 | -1.7362
1.0866 | -33.9571 | -1.7512
1.0927 | -33.9563 | -1.7655
1.0988 | -33.9556 | -1.7803
1.1049 | -33.9548 | -1.7943
1.111 | -33.954 | -1.8097
1.1171 | -33.9533 | -1.8241
1.1232 | -33.9528 | -1.8391
1.1293 | -33.9524 | -1.8531
1.1354 | -33.9521 | -1.8679
1.1415 | -33.9517 | -1.8833
1.1476 | -33.951 | -1.8987
1.1537 | -33.9505 | -1.9132
1.1598 | -33.9506 | -1.9278
1.1659 | -33.9503 | -1.9409
1.172 | -33.95 | -1.9547
1.1781 | -33.9498 | -1.9685
1.1842 | -33.9496 | -1.9826
1.1903 | -33.9499 | -1.9935
1.1964 | -33.9498 | -2.0068
1.2025 | -33.9499 | -2.0206
1.2086 | -33.9502 | -2.0329
1.2147 | -33.9501 | -2.0453
1.2209 | -33.9505 | -2.0564
1.227 | -33.9511 | -2.0671
1.2331 | -33.9515 | -2.079
1.2392 | -33.9518 | -2.0908
1.2453 | -33.9522 | -2.1018
1.2514 | -33.9526 | -2.1118
1.2575 | -33.9531 | -2.1227
1.2636 | -33.9539 | -2.1316
1.2697 | -33.9545 | -2.1408
1.2758 | -33.9551 | -2.1518
1.2819 | -33.9558 | -2.1619
1.288 | -33.9564 | -2.1727
1.2941 | -33.9571 | -2.1812
1.3002 | -33.9578 | -2.1893
1.3063 | -33.9582 | -2.1983
1.3124 | -33.9589 | -2.2073
1.3185 | -33.9595 | -2.2164
1.3246 | -33.9601 | -2.2255
1.3307 | -33.9609 | -2.234
1.3368 | -33.9617 | -2.2426
1.3429 | -33.9622 | -2.2515
1.349 | -33.9627 | -2.2608
1.3551 | -33.9633 | -2.27
1.3613 | -33.9638 | -2.2781
1.3674 | -33.9644 | -2.2879
1.3735 | -33.9648 | -2.2975
1.3796 | -33.9652 | -2.3064
1.3857 | -33.9655 | -2.3155
1.3918 | -33.966 | -2.3256
1.3979 | -33.9665 | -2.3347
1.404 | -33.9669 | -2.3459
1.4101 | -33.9673 | -2.356
1.4162 | -33.9676 | -2.3663
1.4223 | -33.9681 | -2.3775
1.4284 | -33.9685 | -2.3873
1.4345 | -33.9687 | -2.3987
1.4406 | -33.9689 | -2.4121
1.4467 | -33.9693 | -2.4235
1.4528 | -33.9698 | -2.4346
1.4589 | -33.9702 | -2.4463
1.465 | -33.9704 | -2.457
1.4711 | -33.9708 | -2.4682
1.4772 | -33.9713 | -2.4814
1.4833 | -33.9718 | -2.4942
1.4894 | -33.9723 | -2.5062
1.4955 | -33.9728 | -2.5181
1.5016 | -33.9733 | -2.5304
1.5078 | -33.9736 | -2.5442
1.5139 | -33.974 | -2.5572
1.52 | -33.9744 | -2.5696
1.5261 | -33.9747 | -2.5822
1.5322 | -33.975 | -2.5961
1.5383 | -33.9754 | -2.6086
1.5444 | -33.9759 | -2.6213
1.5505 | -33.9764 | -2.6347
1.5566 | -33.9769 | -2.6477
1.5627 | -33.9773 | -2.6615
1.5688 | -33.9779 | -2.675
1.5749 | -33.9782 | -2.6877
1.581 | -33.9785 | -2.701
1.5871 | -33.9789 | -2.7148
1.5932 | -33.9792 | -2.7276
1.5993 | -33.9796 | -2.7418
1.6054 | -33.9799 | -2.755
1.6115 | -33.9804 | -2.7692
1.6176 | -33.9806 | -2.7818
1.6237 | -33.9808 | -2.7955
1.6298 | -33.9811 | -2.8082
1.6359 | -33.9814 | -2.821
1.642 | -33.9818 | -2.833
1.6482 | -33.982 | -2.8461
1.6543 | -33.9823 | -2.8599
1.6604 | -33.9826 | -2.8719
1.6665 | -33.9826 | -2.8845
1.6726 | -33.9828 | -2.8972
1.6787 | -33.9831 | -2.9098
1.6848 | -33.983 | -2.9224
1.6909 | -33.9832 | -2.9352
1.697 | -33.9834 | -2.9474
1.7031 | -33.9834 | -2.9601
1.7092 | -33.9834 | -2.9724
1.7153 | -33.9835 | -2.9845
1.7214 | -33.9832 | -2.9966
1.7275 | -33.9831 | -3.0093
1.7336 | -33.983 | -3.0223
1.7397 | -33.9831 | -3.0347
1.7458 | -33.9831 | -3.047
1.7519 | -33.983 | -3.0592
1.758 | -33.983 | -3.0712
1.7641 | -33.9831 | -3.0828
1.7702 | -33.9831 | -3.0959
1.7763 | -33.983 | -3.1074
1.7824 | -33.9829 | -3.1199
1.7885 | -33.9828 | -3.1323
1.7947 | -33.9829 | -3.1443
1.8008 | -33.9829 | -3.1557
1.8069 | -33.9828 | -3.1673
1.813 | -33.9828 | -3.1796
1.8191 | -33.9828 | -3.1924
1.8252 | -33.9828 | -3.205
1.8313 | -33.9829 | -3.2172
1.8374 | -33.983 | -3.2289
1.8435 | -33.9831 | -3.2408
1.8496 | -33.9831 | -3.2537
1.8557 | -33.9832 | -3.2664
1.8618 | -33.9835 | -3.2785
1.8679 | -33.9835 | -3.2905
1.874 | -33.9836 | -3.3024
1.8801 | -33.9837 | -3.3147
1.8862 | -33.984 | -3.3266
1.8923 | -33.9841 | -3.3386
1.8984 | -33.9842 | -3.3505
1.9045 | -33.9842 | -3.362
1.9106 | -33.9842 | -3.3741
1.9167 | -33.9843 | -3.3862
1.9228 | -33.9844 | -3.3979
1.9289 | -33.9845 | -3.4095
1.9351 | -33.9846 | -3.4202
1.9412 | -33.9848 | -3.4313
1.9473 | -33.9851 | -3.4433
1.9534 | -33.9852 | -3.4545
1.9595 | -33.9854 | -3.4661
1.9656 | -33.9856 | -3.4781
1.9717 | -33.9858 | -3.4896
1.9778 | -33.9861 | -3.5008
1.9839 | -33.9863 | -3.512
1.99 | -33.9865 | -3.5228
1.9961 | -33.9866 | -3.5347
2.0022 | -33.9868 | -3.5458
2.0083 | -33.9869 | -3.5569
2.0144 | -33.9871 | -3.5681
2.0205 | -33.9874 | -3.5797
2.0266 | -33.9877 | -3.591
2.0327 | -33.9879 | -3.6023
2.0388 | -33.9883 | -3.6136
2.0449 | -33.9885 | -3.6247
2.051 | -33.9887 | -3.6358
2.0571 | -33.989 | -3.647
2.0632 | -33.9892 | -3.6582
2.0693 | -33.9893 | -3.6702
2.0754 | -33.9896 | -3.6814
2.0816 | -33.9899 | -3.6934
2.0877 | -33.99 | -3.7047
2.0938 | -33.9904 | -3.7166
2.0999 | -33.9905 | -3.7281
2.106 | -33.9907 | -3.7395
2.1121 | -33.9909 | -3.7514
2.1182 | -33.9911 | -3.762
2.1243 | -33.9913 | -3.7726
2.1304 | -33.9915 | -3.7836
2.1365 | -33.9918 | -3.795
2.1426 | -33.9919 | -3.8069
2.1487 | -33.9922 | -3.818
2.1548 | -33.9926 | -3.8295
2.1609 | -33.993 | -3.8408
2.167 | -33.9934 | -3.8519
2.1731 | -33.9936 | -3.8637
2.1792 | -33.9938 | -3.8746
2.1853 | -33.9942 | -3.8863
2.1914 | -33.9944 | -3.8975
2.1975 | -33.9948 | -3.9088
2.2036 | -33.995 | -3.9207
2.2097 | -33.9952 | -3.932
2.2158 | -33.9956 | -3.9424
2.222 | -33.9958 | -3.9527
2.2281 | -33.996 | -3.9639
2.2342 | -33.9963 | -3.9751
2.2403 | -33.9964 | -3.9856
2.2464 | -33.9967 | -3.9963
2.2525 | -33.9969 | -4.0076
2.2586 | -33.9972 | -4.0181
2.2647 | -33.9975 | -4.0284
2.2708 | -33.9978 | -4.0395
2.2769 | -33.9981 | -4.05
2.283 | -33.9982 | -4.061
2.2891 | -33.9984 | -4.0718
2.2952 | -33.9985 | -4.0822
2.3013 | -33.9987 | -4.0929
2.3074 | -33.999 | -4.1033
2.3135 | -33.9991 | -4.1145
2.3196 | -33.9993 | -4.1245
2.3257 | -33.9995 | -4.135
2.3318 | -33.9996 | -4.1458
2.3379 | -33.9998 | -4.1565
2.344 | -33.9999 | -4.1677
2.3501 | -34.0001 | -4.1784
2.3562 | -34.0002 | -4.1887
2.3623 | -34.0003 | -4.1989
2.3685 | -34.0004 | -4.2093
2.3746 | -34.0005 | -4.2204
2.3807 | -34.0007 | -4.2307
2.3868 | -34.0009 | -4.2412
2.3929 | -34.001 | -4.2517
2.399 | -34.0011 | -4.2621
2.4051 | -34.0013 | -4.2724
2.4112 | -34.0015 | -4.2829
2.4173 | -34.0016 | -4.2933
2.4234 | -34.0017 | -4.3039
2.4295 | -34.0018 | -4.3146
2.4356 | -34.0019 | -4.3254
2.4417 | -34.0021 | -4.3363
2.4478 | -34.0021 | -4.347
2.4539 | -34.0023 | -4.3571
2.46 | -34.0024 | -4.3684
2.4661 | -34.0025 | -4.3794
2.4722 | -34.0027 | -4.3894
2.4783 | -34.0027 | -4.3997
2.4844 | -34.0028 | -4.4106
2.4905 | -34.0029 | -4.4213
2.4966 | -34.0031 | -4.4319
2.5027 | -34.0031 | -4.4422
2.5089 | -34.0032 | -4.4528
2.515 | -34.0032 | -4.4635
2.5211 | -34.0033 | -4.4741
2.5272 | -34.0034 | -4.4848
2.5333 | -34.0035 | -4.4956
2.5394 | -34.0037 | -4.5062
2.5455 | -34.0038 | -4.5174
2.5516 | -34.0039 | -4.5275
2.5577 | -34.0041 | -4.5388
2.5638 | -34.0042 | -4.5503
2.5699 | -34.0042 | -4.5615
2.576 | -34.0043 | -4.5729
2.5821 | -34.0044 | -4.584
2.5882 | -34.0045 | -4.5949
2.5943 | -34.0046 | -4.6055
2.6004 | -34.0048 | -4.6164
2.6065 | -34.0049 | -4.6266
2.6126 | -34.005 | -4.637
2.6187 | -34.0051 | -4.6474
2.6248 | -34.0052 | -4.6581
2.6309 | -34.0054 | -4.6694
2.637 | -34.0056 | -4.6809
2.6431 | -34.0058 | -4.6912
2.6492 | -34.0059 | -4.7024
2.6554 | -34.0061 | -4.7131
2.6615 | -34.0062 | -4.724
2.6676 | -34.0064 | -4.7353
2.6737 | -34.0067 | -4.7464
2.6798 | -34.0069 | -4.7574
2.6859 | -34.0071 | -4.7686
2.692 | -34.0073 | -4.7793
2.6981 | -34.0074 | -4.7904
2.7042 | -34.0076 | -4.8015
2.7103 | -34.0078 | -4.812
2.7164 | -34.0081 | -4.8228
2.7225 | -34.0083 | -4.834
2.7286 | -34.0086 | -4.845
2.7347 | -34.0088 | -4.8554
2.7408 | -34.009 | -4.8655
2.7469 | -34.0092 | -4.8767
2.753 | -34.0095 | -4.8872
2.7591 | -34.0098 | -4.8984
2.7652 | -34.01 | -4.9098
2.7713 | -34.0101 | -4.9205
2.7774 | -34.0104 | -4.9314
2.7835 | -34.0105 | -4.9423
2.7896 | -34.0107 | -4.9528
2.7958 | -34.011 | -4.9636
2.8019 | -34.0113 | -4.974
2.808 | -34.0115 | -4.9847
2.8141 | -34.0118 | -4.9958
2.8202 | -34.0121 | -5.0067
2.8263 | -34.0124 | -5.0176
2.8324 | -34.0128 | -5.0289
2.8385 | -34.0131 | -5.0397
2.8446 | -34.0135 | -5.0505
2.8507 | -34.0138 | -5.0614
2.8568 | -34.0141 | -5.0718
2.8629 | -34.0144 | -5.0824
2.869 | -34.0147 | -5.0936
2.8751 | -34.015 | -5.1042
2.8812 | -34.0153 | -5.1155
2.8873 | -34.0156 | -5.1264
2.8934 | -34.0159 | -5.1374
2.8995 | -34.0162 | -5.1485
2.9056 | -34.0165 | -5.1597
2.9117 | -34.0168 | -5.1707
2.9178 | -34.0171 | -5.1816
2.9239 | -34.0175 | -5.1931
2.93 | -34.0179 | -5.2046
2.9361 | -34.0181 | -5.2153
2.9423 | -34.0185 | -5.226
2.9484 | -34.0188 | -5.2371
2.9545 | -34.0192 | -5.2482
2.9606 | -34.0196 | -5.2597
2.9667 | -34.0199 | -5.2708
2.9728 | -34.0202 | -5.282
2.9789 | -34.0206 | -5.2932
2.985 | -34.021 | -5.3045
2.9911 | -34.0213 | -5.3158
2.9972 | -34.0217 | -5.3268
3.0033 | -34.0221 | -5.3384
3.0094 | -34.0225 | -5.3495
3.0155 | -34.0229 | -5.3604
3.0216 | -34.0232 | -5.3717
3.0277 | -34.0235 | -5.3834
3.0338 | -34.0239 | -5.3948
3.0399 | -34.0242 | -5.4062
3.046 | -34.0245 | -5.4178
3.0521 | -34.0248 | -5.4289
3.0582 | -34.0251 | -5.4408
3.0643 | -34.0254 | -5.4523
3.0704 | -34.0257 | -5.4641
3.0765 | -34.0259 | -5.4753
3.0827 | -34.0262 | -5.4864
3.0888 | -34.0265 | -5.4977
3.0949 | -34.0269 | -5.5092
3.101 | -34.0272 | -5.5206
3.1071 | -34.0275 | -5.5324
3.1132 | -34.0278 | -5.5442
3.1193 | -34.0281 | -5.5556
3.1254 | -34.0285 | -5.567
3.1315 | -34.0288 | -5.5789
3.1376 | -34.0291 | -5.5902
3.1437 | -34.0294 | -5.6018
3.1498 | -34.0297 | -5.6138
3.1559 | -34.03 | -5.6253
3.162 | -34.0302 | -5.6373
3.1681 | -34.0304 | -5.6492
3.1742 | -34.0307 | -5.6608
3.1803 | -34.031 | -5.6724
3.1864 | -34.0313 | -5.6837
3.1925 | -34.0316 | -5.6952
3.1986 | -34.0318 | -5.7071
3.2047 | -34.0321 | -5.719
3.2108 | -34.0323 | -5.7301
3.2169 | -34.0326 | -5.7417
3.223 | -34.0328 | -5.7533
3.2292 | -34.033 | -5.765
3.2353 | -34.0332 | -5.7757
3.2414 | -34.0334 | -5.7871
3.2475 | -34.0336 | -5.7985
3.2536 | -34.0339 | -5.8099
3.2597 | -34.0341 | -5.8208
3.2658 | -34.0343 | -5.8317
3.2719 | -34.0345 | -5.8427
3.278 | -34.0347 | -5.8543
3.2841 | -34.0349 | -5.8653
3.2902 | -34.0351 | -5.8766
3.2963 | -34.0354 | -5.8881
3.3024 | -34.0355 | -5.8991
3.3085 | -34.0357 | -5.9106
3.3146 | -34.0359 | -5.9218
3.3207 | -34.0361 | -5.9332
3.3268 | -34.0363 | -5.9446
3.3329 | -34.0365 | -5.9552
3.339 | -34.0366 | -5.966
3.3451 | -34.0368 | -5.9771
3.3512 | -34.0369 | -5.9877
3.3573 | -34.0371 | -5.9989
3.3634 | -34.0373 | -6.0099
3.3696 | -34.0375 | -6.0213
3.3757 | -34.0376 | -6.032
3.3818 | -34.0379 | -6.0429
3.3879 | -34.038 | -6.0539
3.394 | -34.0381 | -6.0653
3.4001 | -34.0383 | -6.0761
3.4062 | -34.0384 | -6.087
3.4123 | -34.0385 | -6.0975
3.4184 | -34.0387 | -6.1088
3.4245 | -34.039 | -6.1195
3.4306 | -34.039 | -6.1303
3.4367 | -34.0392 | -6.1408
3.4428 | -34.0394 | -6.1511
3.4489 | -34.0395 | -6.162
3.455 | -34.0396 | -6.1728
3.4611 | -34.0398 | -6.1841
3.4672 | -34.04 | -6.195
3.4733 | -34.0401 | -6.206
3.4794 | -34.0403 | -6.2168
3.4855 | -34.0405 | -6.2272
3.4916 | -34.0406 | -6.2376
3.4977 | -34.0408 | -6.2488
3.5038 | -34.0409 | -6.2597
3.5099 | -34.041 | -6.27
3.5161 | -34.0412 | -6.2804
3.5222 | -34.0413 | -6.2907
3.5283 | -34.0414 | -6.3014
3.5344 | -34.0415 | -6.3121
3.5405 | -34.0417 | -6.3227
3.5466 | -34.0419 | -6.3336
3.5527 | -34.042 | -6.3443
3.5588 | -34.0421 | -6.3548
3.5649 | -34.0421 | -6.3656
3.571 | -34.0423 | -6.3761
3.5771 | -34.0424 | -6.3872
3.5832 | -34.0425 | -6.3971
3.5893 | -34.0428 | -6.4076
3.5954 | -34.0429 | -6.4184
3.6015 | -34.0431 | -6.4293
3.6076 | -34.0433 | -6.4398
3.6137 | -34.0434 | -6.4498
3.6198 | -34.0437 | -6.46
3.6259 | -34.0438 | -6.4709
3.632 | -34.044 | -6.4815
3.6381 | -34.0442 | -6.4922
3.6442 | -34.0443 | -6.5027
3.6503 | -34.0445 | -6.5129
3.6565 | -34.0447 | -6.5236
3.6626 | -34.0448 | -6.534
3.6687 | -34.0451 | -6.5446
3.6748 | -34.0453 | -6.5548
3.6809 | -34.0456 | -6.5654
3.687 | -34.0457 | -6.5764
3.6931 | -34.046 | -6.5864
3.6992 | -34.0462 | -6.5971
3.7053 | -34.0464 | -6.6074
3.7114 | -34.0465 | -6.6181
3.7175 | -34.0467 | -6.6288
3.7236 | -34.047 | -6.6388
3.7297 | -34.0472 | -6.6495
3.7358 | -34.0473 | -6.6603
3.7419 | -34.0475 | -6.671
3.748 | -34.0478 | -6.6815
3.7541 | -34.0479 | -6.692
3.7602 | -34.0482 | -6.7024
3.7663 | -34.0484 | -6.7128
3.7724 | -34.0486 | -6.7234
3.7785 | -34.0489 | -6.7337
3.7846 | -34.0491 | -6.7441
3.7907 | -34.0493 | -6.755
3.7969 | -34.0496 | -6.7656
3.803 | -34.0498 | -6.7758
3.8091 | -34.05 | -6.7862
3.8152 | -34.0503 | -6.7967
3.8213 | -34.0505 | -6.8075
3.8274 | -34.0507 | -6.8178
3.8335 | -34.051 | -6.8285
3.8396 | -34.0513 | -6.8392
3.8457 | -34.0516 | -6.8498
3.8518 | -34.0519 | -6.8602
3.8579 | -34.0522 | -6.8706
3.864 | -34.0524 | -6.8813
3.8701 | -34.0527 | -6.8922
3.8762 | -34.0529 | -6.9025
3.8823 | -34.0531 | -6.9133
3.8884 | -34.0534 | -6.9238
3.8945 | -34.0537 | -6.9346
3.9006 | -34.054 | -6.9455
3.9067 | -34.0542 | -6.9566
3.9128 | -34.0545 | -6.9672
3.9189 | -34.0547 | -6.9776
3.925 | -34.0549 | -6.9884
3.9311 | -34.0552 | -6.9986
3.9372 | -34.0555 | -7.0094
3.9434 | -34.0557 | -7.0201
3.9495 | -34.0561 | -7.0308
3.9556 | -34.0564 | -7.0413
3.9617 | -34.0566 | -7.0515
3.9678 | -34.0569 | -7.0619
3.9739 | -34.0571 | -7.0725
3.98 | -34.0574 | -7.0831
3.9861 | -34.0576 | -7.0938
3.9922 | -34.0578 | -7.1044
3.9983 | -34.0581 | -7.1147
4.0044 | -34.0583 | -7.1249
4.0105 | -34.0586 | -7.1357
4.0166 | -34.0589 | -7.1465
4.0227 | -34.0591 | -7.157
4.0288 | -34.0593 | -7.168
4.0349 | -34.0595 | -7.1784
4.041 | -34.0597 | -7.1891
4.0471 | -34.0599 | -7.1998
4.0532 | -34.0602 | -7.2101
4.0593 | -34.0605 | -7.2209
4.0654 | -34.0607 | -7.2313
4.0715 | -34.061 | -7.2424
4.0776 | -34.0613 | -7.253
4.0838 | -34.0615 | -7.264
4.0899 | -34.0618 | -7.2743
4.096 | -34.062 | -7.2849
4.1021 | -34.0622 | -7.2953
4.1082 | -34.0625 | -7.3057
4.1143 | -34.0627 | -7.316
4.1204 | -34.063 | -7.3265
4.1265 | -34.0632 | -7.3368
4.1326 | -34.0634 | -7.3472
4.1387 | -34.0637 | -7.3577
4.1448 | -34.0639 | -7.3683
4.1509 | -34.0641 | -7.379
4.157 | -34.0643 | -7.39
4.1631 | -34.0646 | -7.4007
4.1692 | -34.0648 | -7.4113
4.1753 | -34.0651 | -7.4219
4.1814 | -34.0652 | -7.4322
4.1875 | -34.0655 | -7.4431
4.1936 | -34.0657 | -7.4538
4.1997 | -34.0659 | -7.4639
4.2058 | -34.0661 | -7.4748
4.2119 | -34.0664 | -7.4852
4.218 | -34.0666 | -7.4955
4.2241 | -34.0668 | -7.5058
4.2303 | -34.067 | -7.5162
4.2364 | -34.0672 | -7.5268
4.2425 | -34.0674 | -7.5374
4.2486 | -34.0676 | -7.5481
4.2547 | -34.0677 | -7.5584
4.2608 | -34.0679 | -7.5691
4.2669 | -34.0681 | -7.5793
4.273 | -34.0683 | -7.59
4.2791 | -34.0685 | -7.6005
4.2852 | -34.0687 | -7.6111
4.2913 | -34.0689 | -7.6213
4.2974 | -34.0691 | -7.6322
4.3035 | -34.0694 | -7.643
4.3096 | -34.0695 | -7.6537
4.3157 | -34.0697 | -7.6641
4.3218 | -34.0699 | -7.6745
4.3279 | -34.0701 | -7.6849
4.334 | -34.0703 | -7.6953
4.3401 | -34.0705 | -7.7054
4.3462 | -34.0707 | -7.7161
4.3523 | -34.071 | -7.7261
4.3584 | -34.0711 | -7.7362
4.3645 | -34.0713 | -7.7467
4.3707 | -34.0715 | -7.7568
4.3768 | -34.0717 | -7.7672
4.3829 | -34.072 | -7.778
4.389 | -34.0721 | -7.7884
4.3951 | -34.0724 | -7.7985
4.4012 | -34.0725 | -7.8092
4.4073 | -34.0727 | -7.8193
4.4134 | -34.0729 | -7.8296
4.4195 | -34.0732 | -7.8399
4.4256 | -34.0734 | -7.8501
4.4317 | -34.0736 | -7.8604
4.4378 | -34.0738 | -7.8704
4.4439 | -34.074 | -7.8805
4.45 | -34.0741 | -7.8907
4.4561 | -34.0743 | -7.901
4.4622 | -34.0745 | -7.9113
4.4683 | -34.0747 | -7.9215
4.4744 | -34.0749 | -7.9322
4.4805 | -34.0751 | -7.942
4.4866 | -34.0753 | -7.9523
4.4927 | -34.0755 | -7.9628
4.4988 | -34.0757 | -7.973
4.5049 | -34.0759 | -7.9829
4.511 | -34.0762 | -7.9928
4.5172 | -34.0764 | -8.0031
4.5233 | -34.0766 | -8.0137
4.5294 | -34.0768 | -8.0238
4.5355 | -34.077 | -8.034
4.5416 | -34.0772 | -8.0441
4.5477 | -34.0775 | -8.0543
4.5538 | -34.0776 | -8.0642
4.5599 | -34.0779 | -8.0742
4.566 | -34.0781 | -8.0844
4.5721 | -34.0783 | -8.0945
4.5782 | -34.0786 | -8.1048
4.5843 | -34.0788 | -8.1151
4.5904 | -34.0791 | -8.1254
4.5965 | -34.0793 | -8.1354
4.6026 | -34.0795 | -8.146
4.6087 | -34.0797 | -8.156
4.6148 | -34.0799 | -8.1665
4.6209 | -34.0802 | -8.1766
4.627 | -34.0804 | -8.187
4.6331 | -34.0806 | -8.1969
4.6392 | -34.0808 | -8.207
4.6453 | -34.0811 | -8.2171
4.6514 | -34.0813 | -8.2273
4.6576 | -34.0815 | -8.2378
4.6637 | -34.0817 | -8.2476
4.6698 | -34.0819 | -8.2575
4.6759 | -34.0821 | -8.268
4.682 | -34.0823 | -8.2781
4.6881 | -34.0825 | -8.2881
4.6942 | -34.0827 | -8.2982
4.7003 | -34.0829 | -8.3083
4.7064 | -34.0832 | -8.3183
4.7125 | -34.0834 | -8.3286
4.7186 | -34.0836 | -8.3384
4.7247 | -34.0839 | -8.3487
4.7308 | -34.0841 | -8.3587
4.7369 | -34.0843 | -8.3689
4.743 | -34.0846 | -8.3795
4.7491 | -34.0848 | -8.3898
4.7552 | -34.085 | -8.4
4.7613 | -34.0852 | -8.4102
4.7674 | -34.0854 | -8.4202
4.7735 | -34.0856 | -8.4307
4.7796 | -34.0859 | -8.441
4.7857 | -34.0861 | -8.4512
4.7918 | -34.0864 | -8.4612
4.7979 | -34.0866 | -8.4712
4.8041 | -34.0869 | -8.4812
4.8102 | -34.0871 | -8.4911
4.8163 | -34.0873 | -8.501
4.8224 | -34.0876 | -8.5112
4.8285 | -34.0878 | -8.5209
4.8346 | -34.0881 | -8.5308
4.8407 | -34.0883 | -8.5411
4.8468 | -34.0886 | -8.5514
4.8529 | -34.0888 | -8.5617
4.859 | -34.089 | -8.5722
4.8651 | -34.0892 | -8.5821
4.8712 | -34.0895 | -8.5919
4.8773 | -34.0896 | -8.6021
4.8834 | -34.0899 | -8.6122
4.8895 | -34.0901 | -8.622
4.8956 | -34.0903 | -8.6322
4.9017 | -34.0905 | -8.6422
4.9078 | -34.0907 | -8.6526
4.9139 | -34.091 | -8.6623
4.92 | -34.0913 | -8.6721
4.9261 | -34.0916 | -8.6818
4.9322 | -34.0918 | -8.6919
4.9383 | -34.0921 | -8.702
4.9445 | -34.0923 | -8.7122
4.9506 | -34.0925 | -8.7221
4.9567 | -34.0927 | -8.732
4.9628 | -34.093 | -8.7421
4.9689 | -34.0932 | -8.7525
4.975 | -34.0934 | -8.7629
4.9811 | -34.0937 | -8.7727
4.9872 | -34.094 | -8.7827
4.9933 | -34.0943 | -8.7931
4.9994 | -34.0945 | -8.8036
5.0055 | -34.0948 | -8.8137
5.0116 | -34.095 | -8.8242
5.0177 | -34.0953 | -8.8346
5.0238 | -34.0956 | -8.8451
5.0299 | -34.0959 | -8.8556
5.036 | -34.0962 | -8.8658
5.0421 | -34.0964 | -8.8758
5.0482 | -34.0967 | -8.8859
5.0543 | -34.0969 | -8.8958
5.0604 | -34.0972 | -8.9061
5.0665 | -34.0975 | -8.9161
5.0726 | -34.0977 | -8.9266
5.0787 | -34.098 | -8.9369
5.0848 | -34.0983 | -8.9473
5.091 | -34.0986 | -8.9573
5.0971 | -34.0989 | -8.9672
5.1032 | -34.0991 | -8.9777
5.1093 | -34.0994 | -8.9879
5.1154 | -34.0996 | -8.9979
5.1215 | -34.0998 | -9.0079
5.1276 | -34.1001 | -9.0178
5.1337 | -34.1004 | -9.0278
5.1398 | -34.1007 | -9.0376
5.1459 | -34.101 | -9.0481
5.152 | -34.1012 | -9.0585
5.1581 | -34.1015 | -9.0689
5.1642 | -34.1018 | -9.0791
5.1703 | -34.102 | -9.0893
5.1764 | -34.1023 | -9.0994
5.1825 | -34.1026 | -9.1094
5.1886 | -34.1029 | -9.1193
5.1947 | -34.1031 | -9.1295
5.2008 | -34.1034 | -9.1399
5.2069 | -34.1037 | -9.1503
5.213 | -34.104 | -9.1603
5.2191 | -34.1043 | -9.1706
5.2252 | -34.1045 | -9.181
5.2314 | -34.1048 | -9.1914
5.2375 | -34.105 | -9.2015
5.2436 | -34.1052 | -9.2119
5.2497 | -34.1055 | -9.2218
5.2558 | -34.1058 | -9.2325
5.2619 | -34.1061 | -9.243
5.268 | -34.1064 | -9.2535
5.2741 | -34.1067 | -9.2639
5.2802 | -34.107 | -9.2746
5.2863 | -34.1073 | -9.2848
5.2924 | -34.1076 | -9.2952
5.2985 | -34.1079 | -9.3056
5.3046 | -34.1082 | -9.3159
5.3107 | -34.1085 | -9.3266
5.3168 | -34.1087 | -9.337
5.3229 | -34.109 | -9.3473
5.329 | -34.1093 | -9.3577
5.3351 | -34.1096 | -9.3677
5.3412 | -34.1098 | -9.378
5.3473 | -34.1101 | -9.3885
5.3534 | -34.1104 | -9.399
5.3595 | -34.1107 | -9.4093
5.3656 | -34.111 | -9.4196
5.3717 | -34.1112 | -9.4302
5.3779 | -34.1115 | -9.4411
5.384 | -34.1118 | -9.4514
5.3901 | -34.1121 | -9.4621
5.3962 | -34.1124 | -9.4723
5.4023 | -34.1127 | -9.4829
5.4084 | -34.113 | -9.4931
5.4145 | -34.1133 | -9.5034
5.4206 | -34.1136 | -9.5138
5.4267 | -34.1139 | -9.5246
5.4328 | -34.1142 | -9.5349
5.4389 | -34.1144 | -9.5451
5.445 | -34.1147 | -9.5558
5.4511 | -34.115 | -9.5662
5.4572 | -34.1153 | -9.5767
5.4633 | -34.1156 | -9.5871
5.4694 | -34.1159 | -9.5975
5.4755 | -34.1162 | -9.6076
5.4816 | -34.1164 | -9.6181
5.4877 | -34.1168 | -9.6284
5.4938 | -34.117 | -9.6387
5.4999 | -34.1174 | -9.6489
5.506 | -34.1177 | -9.6594
5.5121 | -34.1179 | -9.6698
5.5183 | -34.1182 | -9.6799
5.5244 | -34.1185 | -9.6906
5.5305 | -34.1188 | -9.7009
5.5366 | -34.1191 | -9.7111
5.5427 | -34.1194 | -9.7217
5.5488 | -34.1197 | -9.7322
5.5549 | -34.1199 | -9.7427
5.561 | -34.1202 | -9.7534
5.5671 | -34.1205 | -9.7637
5.5732 | -34.1208 | -9.7744
5.5793 | -34.1211 | -9.7852
5.5854 | -34.1215 | -9.7955
5.5915 | -34.1217 | -9.8064
5.5976 | -34.122 | -9.8168
5.6037 | -34.1223 | -9.8273
5.6098 | -34.1226 | -9.8377
5.6159 | -34.1229 | -9.8479
5.622 | -34.1231 | -9.8584
5.6281 | -34.1235 | -9.8689
5.6342 | -34.1238 | -9.8797
5.6403 | -34.124 | -9.8902
5.6464 | -34.1243 | -9.9008
5.6525 | -34.1246 | -9.9113
5.6586 | -34.1249 | -9.9215
5.6648 | -34.1252 | -9.932
5.6709 | -34.1256 | -9.9427
5.677 | -34.1258 | -9.9532
5.6831 | -34.1261 | -9.9638
5.6892 | -34.1264 | -9.9744
5.6953 | -34.1267 | -9.9849
5.7014 | -34.127 | -9.9954
5.7075 | -34.1273 | -10.0057
5.7136 | -34.1275 | -10.0161
5.7197 | -34.1278 | -10.0265
5.7258 | -34.1281 | -10.0368
5.7319 | -34.1284 | -10.0473
5.738 | -34.1287 | -10.0578
5.7441 | -34.129 | -10.0682
5.7502 | -34.1293 | -10.0785
5.7563 | -34.1296 | -10.089
5.7624 | -34.1299 | -10.0997
5.7685 | -34.1302 | -10.1102
5.7746 | -34.1305 | -10.1209
5.7807 | -34.1307 | -10.1314
5.7868 | -34.131 | -10.1421
5.7929 | -34.1312 | -10.1525
5.799 | -34.1315 | -10.1627
5.8052 | -34.1318 | -10.1731
5.8113 | -34.1321 | -10.1833
5.8174 | -34.1324 | -10.1937
5.8235 | -34.1327 | -10.2041
5.8296 | -34.133 | -10.2148
5.8357 | -34.1332 | -10.2251
5.8418 | -34.1335 | -10.2357
5.8479 | -34.1338 | -10.2462
5.854 | -34.1341 | -10.2569
5.8601 | -34.1344 | -10.2672
5.8662 | -34.1346 | -10.2778
5.8723 | -34.1349 | -10.2881
5.8784 | -34.1352 | -10.2984
5.8845 | -34.1355 | -10.3088
5.8906 | -34.1357 | -10.3193
5.8967 | -34.136 | -10.3296
5.9028 | -34.1363 | -10.3398
5.9089 | -34.1367 | -10.3501
5.915 | -34.1369 | -10.3603
5.9211 | -34.1372 | -10.3707
5.9272 | -34.1375 | -10.3809
5.9333 | -34.1378 | -10.3911
5.9394 | -34.1381 | -10.4014
5.9455 | -34.1384 | -10.4116
5.9517 | -34.1387 | -10.4218
5.9578 | -34.1391 | -10.4324
5.9639 | -34.1394 | -10.4424
5.97 | -34.1397 | -10.4529
5.9761 | -34.14 | -10.4631
5.9822 | -34.1403 | -10.4736
5.9883 | -34.1406 | -10.484
5.9944 | -34.1409 | -10.4944
6.0005 | -34.1412 | -10.5049
6.0066 | -34.1415 | -10.5151
6.0127 | -34.1418 | -10.5251
6.0188 | -34.1421 | -10.5357
6.0249 | -34.1424 | -10.5461
6.031 | -34.1426 | -10.5562
6.0371 | -34.1429 | -10.5666
6.0432 | -34.1432 | -10.5768
6.0493 | -34.1435 | -10.5869
6.0554 | -34.1438 | -10.5971
6.0615 | -34.1441 | -10.6074
6.0676 | -34.1444 | -10.6179
6.0737 | -34.1448 | -10.6284
6.0798 | -34.145 | -10.6387
6.0859 | -34.1453 | -10.6489
6.0921 | -34.1456 | -10.659
6.0982 | -34.1459 | -10.6695
6.1043 | -34.1462 | -10.6797
6.1104 | -34.1466 | -10.6902
6.1165 | -34.1469 | -10.7
6.1226 | -34.1472 | -10.7102
6.1287 | -34.1475 | -10.7205
6.1348 | -34.1478 | -10.7308
6.1409 | -34.1481 | -10.741
6.147 | -34.1484 | -10.7512
6.1531 | -34.1487 | -10.7614
6.1592 | -34.149 | -10.7716
6.1653 | -34.1493 | -10.7817
6.1714 | -34.1495 | -10.7921
6.1775 | -34.1499 | -10.8024
6.1836 | -34.1502 | -10.8127
6.1897 | -34.1505 | -10.8229
6.1958 | -34.1508 | -10.8334
6.2019 | -34.1511 | -10.8437
6.208 | -34.1514 | -10.854
6.2141 | -34.1517 | -10.8641
6.2202 | -34.1521 | -10.8745
6.2263 | -34.1524 | -10.8846
6.2325 | -34.1527 | -10.8945
6.2386 | -34.153 | -10.9044
6.2447 | -34.1533 | -10.9149
6.2508 | -34.1536 | -10.9252
6.2569 | -34.1539 | -10.9353
6.263 | -34.1542 | -10.9455
6.2691 | -34.1545 | -10.9558
6.2752 | -34.1548 | -10.966
6.2813 | -34.155 | -10.9762
6.2874 | -34.1554 | -10.9863
6.2935 | -34.1557 | -10.9964
6.2996 | -34.156 | -11.0065
6.3057 | -34.1563 | -11.0168
6.3118 | -34.1565 | -11.027
6.3179 | -34.1568 | -11.0373
6.324 | -34.1572 | -11.0474
6.3301 | -34.1575 | -11.0576
6.3362 | -34.1578 | -11.0677
6.3423 | -34.1581 | -11.0781
6.3484 | -34.1584 | -11.0883
6.3545 | -34.1587 | -11.0983
6.3606 | -34.159 | -11.1081
6.3667 | -34.1593 | -11.1182
6.3728 | -34.1596 | -11.1284
6.379 | -34.1599 | -11.1386
6.3851 | -34.1602 | -11.1485
6.3912 | -34.1605 | -11.1586
6.3973 | -34.1608 | -11.1689
6.4034 | -34.1611 | -11.1788
6.4095 | -34.1614 | -11.1889
6.4156 | -34.1617 | -11.1991
6.4217 | -34.1619 | -11.2091
6.4278 | -34.1623 | -11.2192
6.4339 | -34.1626 | -11.2292
6.44 | -34.1629 | -11.2392
6.4461 | -34.1632 | -11.2495
6.4522 | -34.1635 | -11.2594
6.4583 | -34.1638 | -11.2696
6.4644 | -34.1641 | -11.2798
6.4705 | -34.1644 | -11.2896
6.4766 | -34.1647 | -11.2997
6.4827 | -34.165 | -11.3099
6.4888 | -34.1652 | -11.3198
6.4949 | -34.1656 | -11.3297
6.501 | -34.1658 | -11.34
6.5071 | -34.1662 | -11.3497
6.5132 | -34.1665 | -11.3599
6.5194 | -34.1668 | -11.3701
6.5255 | -34.1671 | -11.3799
6.5316 | -34.1674 | -11.3901
6.5377 | -34.1677 | -11.4002
6.5438 | -34.1679 | -11.41
6.5499 | -34.1683 | -11.4201
6.556 | -34.1686 | -11.4301
6.5621 | -34.1688 | -11.4405
6.5682 | -34.1692 | -11.4506
6.5743 | -34.1695 | -11.4606
6.5804 | -34.1698 | -11.4708
6.5865 | -34.1701 | -11.4809
6.5926 | -34.1704 | -11.4908
6.5987 | -34.1707 | -11.5007
6.6048 | -34.171 | -11.5109
6.6109 | -34.1714 | -11.5209
6.617 | -34.1716 | -11.5311
6.6231 | -34.1719 | -11.5412
6.6292 | -34.1722 | -11.5513
6.6353 | -34.1725 | -11.5616
6.6414 | -34.1729 | -11.5716
6.6475 | -34.1732 | -11.5816
6.6536 | -34.1735 | -11.5916
6.6597 | -34.1738 | -11.6014
6.6659 | -34.1741 | -11.6115
6.672 | -34.1744 | -11.6215
6.6781 | -34.1747 | -11.6316
6.6842 | -34.175 | -11.6415
6.6903 | -34.1754 | -11.6514
6.6964 | -34.1757 | -11.6614
6.7025 | -34.176 | -11.6712
6.7086 | -34.1763 | -11.6813
6.7147 | -34.1766 | -11.6913
6.7208 | -34.1769 | -11.7015
6.7269 | -34.1772 | -11.7114
6.733 | -34.1775 | -11.7213
6.7391 | -34.1778 | -11.7313
6.7452 | -34.1782 | -11.741
6.7513 | -34.1785 | -11.7507
6.7574 | -34.1788 | -11.7608
6.7635 | -34.1791 | -11.7706
6.7696 | -34.1794 | -11.7804
6.7757 | -34.1798 | -11.7904
6.7818 | -34.1801 | -11.8004
6.7879 | -34.1804 | -11.8103
6.794 | -34.1807 | -11.8202
6.8001 | -34.181 | -11.8302
6.8063 | -34.1813 | -11.8401
6.8124 | -34.1816 | -11.8503
6.8185 | -34.1819 | -11.8604
6.8246 | -34.1822 | -11.8703
6.8307 | -34.1825 | -11.8802
6.8368 | -34.1828 | -11.8904
6.8429 | -34.1832 | -11.9005
6.849 | -34.1835 | -11.9105
6.8551 | -34.1838 | -11.9204
6.8612 | -34.1841 | -11.9304
6.8673 | -34.1844 | -11.9403
6.8734 | -34.1848 | -11.9504
6.8795 | -34.1851 | -11.9604
6.8856 | -34.1854 | -11.9704
6.8917 | -34.1857 | -11.9803
6.8978 | -34.186 | -11.9901
6.9039 | -34.1864 | -11.9998
6.91 | -34.1867 | -12.0099
6.9161 | -34.187 | -12.0196
6.9222 | -34.1873 | -12.0296
6.9283 | -34.1877 | -12.0393
6.9344 | -34.188 | -12.0493
6.9405 | -34.1884 | -12.0591
6.9466 | -34.1887 | -12.0693
6.9528 | -34.189 | -12.0793
6.9589 | -34.1894 | -12.0894
6.965 | -34.1897 | -12.0994
6.9711 | -34.1901 | -12.1094
6.9772 | -34.1905 | -12.1194
6.9833 | -34.1908 | -12.1293
6.9894 | -34.1911 | -12.1395
6.9955 | -34.1914 | -12.1492
7.0016 | -34.1918 | -12.1591
7.0077 | -34.1921 | -12.169
7.0138 | -34.1924 | -12.1789
7.0199 | -34.1927 | -12.1888
7.026 | -34.1931 | -12.1987
7.0321 | -34.1934 | -12.2084
7.0382 | -34.1937 | -12.218
7.0443 | -34.1941 | -12.2281
7.0504 | -34.1944 | -12.238
7.0565 | -34.1947 | -12.248
7.0626 | -34.1951 | -12.2576
7.0687 | -34.1954 | -12.2672
7.0748 | -34.1958 | -12.2769
7.0809 | -34.1961 | -12.2867
7.087 | -34.1965 | -12.2968
7.0932 | -34.1968 | -12.3067
7.0993 | -34.1971 | -12.3165
7.1054 | -34.1975 | -12.3263
7.1115 | -34.1978 | -12.336
7.1176 | -34.1981 | -12.3458
7.1237 | -34.1984 | -12.3557
7.1298 | -34.1988 | -12.3654
7.1359 | -34.1991 | -12.3753
7.142 | -34.1995 | -12.3849
7.1481 | -34.1998 | -12.3946
7.1542 | -34.2001 | -12.4045
7.1603 | -34.2004 | -12.4143
7.1664 | -34.2008 | -12.4241
7.1725 | -34.2012 | -12.434
7.1786 | -34.2015 | -12.4438
7.1847 | -34.2018 | -12.4537
7.1908 | -34.2022 | -12.4637
7.1969 | -34.2025 | -12.4736
7.203 | -34.2028 | -12.4833
7.2091 | -34.2032 | -12.4929
7.2152 | -34.2035 | -12.5027
7.2213 | -34.2039 | -12.5123
7.2274 | -34.2042 | -12.5221
7.2335 | -34.2045 | -12.5319
7.2397 | -34.2049 | -12.5416
7.2458 | -34.2053 | -12.5512
7.2519 | -34.2056 | -12.561
7.258 | -34.2059 | -12.571
7.2641 | -34.2063 | -12.5809
7.2702 | -34.2066 | -12.5908
7.2763 | -34.207 | -12.6006
7.2824 | -34.2073 | -12.6104
7.2885 | -34.2077 | -12.6203
7.2946 | -34.208 | -12.6301
7.3007 | -34.2083 | -12.64
7.3068 | -34.2087 | -12.65
7.3129 | -34.209 | -12.6598
7.319 | -34.2094 | -12.6697
7.3251 | -34.2097 | -12.6793
7.3312 | -34.2101 | -12.6891
7.3373 | -34.2104 | -12.6986
7.3434 | -34.2108 | -12.7083
7.3495 | -34.2112 | -12.718
7.3556 | -34.2115 | -12.7278
7.3617 | -34.2119 | -12.7378
7.3678 | -34.2122 | -12.7475
7.3739 | -34.2126 | -12.7574
7.3801 | -34.2129 | -12.7674
7.3862 | -34.2133 | -12.7771
7.3923 | -34.2136 | -12.7871
7.3984 | -34.214 | -12.7966
7.4045 | -34.2144 | -12.8062
7.4106 | -34.2147 | -12.816
7.4167 | -34.215 | -12.8259
7.4228 | -34.2154 | -12.8358
7.4289 | -34.2158 | -12.8455
7.435 | -34.2162 | -12.855
7.4411 | -34.2165 | -12.8648
7.4472 | -34.2169 | -12.8747
7.4533 | -34.2172 | -12.8842
7.4594 | -34.2176 | -12.8941
7.4655 | -34.218 | -12.9037
7.4716 | -34.2183 | -12.9137
7.4777 | -34.2186 | -12.9234
7.4838 | -34.219 | -12.9329
7.4899 | -34.2194 | -12.9425
7.496 | -34.2197 | -12.9521
7.5021 | -34.2201 | -12.9618
7.5082 | -34.2204 | -12.9715
7.5143 | -34.2208 | -12.9813
7.5204 | -34.2211 | -12.9911
7.5266 | -34.2215 | -13.0007
7.5327 | -34.2218 | -13.0104
7.5388 | -34.2221 | -13.0204
7.5449 | -34.2225 | -13.0303
7.551 | -34.2229 | -13.0399
7.5571 | -34.2232 | -13.0498
7.5632 | -34.2236 | -13.0595
7.5693 | -34.224 | -13.0691
7.5754 | -34.2243 | -13.0787
7.5815 | -34.2247 | -13.0884
7.5876 | -34.225 | -13.0984
7.5937 | -34.2254 | -13.1083
7.5998 | -34.2258 | -13.118
7.6059 | -34.2262 | -13.128
7.612 | -34.2265 | -13.1377
7.6181 | -34.2269 | -13.1474
7.6242 | -34.2273 | -13.1574
7.6303 | -34.2276 | -13.1674
7.6364 | -34.228 | -13.1773
7.6425 | -34.2283 | -13.1871
7.6486 | -34.2286 | -13.1969
7.6547 | -34.229 | -13.2065
7.6608 | -34.2294 | -13.2163
7.667 | -34.2297 | -13.2261
7.6731 | -34.2301 | -13.2359
7.6792 | -34.2304 | -13.246
7.6853 | -34.2308 | -13.2558
7.6914 | -34.2311 | -13.2657
7.6975 | -34.2315 | -13.2755
7.7036 | -34.2319 | -13.2854
7.7097 | -34.2322 | -13.2953
7.7158 | -34.2326 | -13.3052
7.7219 | -34.233 | -13.3147
7.728 | -34.2333 | -13.3246
7.7341 | -34.2337 | -13.3343
7.7402 | -34.2341 | -13.3442
7.7463 | -34.2344 | -13.354
7.7524 | -34.2348 | -13.3639
7.7585 | -34.2351 | -13.3737
7.7646 | -34.2355 | -13.3836
7.7707 | -34.2358 | -13.3934
7.7768 | -34.2361 | -13.4033
7.7829 | -34.2365 | -13.4134
7.789 | -34.2368 | -13.4234
7.7951 | -34.2372 | -13.4332
7.8012 | -34.2375 | -13.4431
7.8073 | -34.2379 | -13.4531
7.8135 | -34.2383 | -13.4629
7.8196 | -34.2387 | -13.4729
7.8257 | -34.239 | -13.4829
7.8318 | -34.2394 | -13.4928
7.8379 | -34.2398 | -13.5025
7.844 | -34.2401 | -13.5123
7.8501 | -34.2405 | -13.5223
7.8562 | -34.2408 | -13.5324
7.8623 | -34.2412 | -13.5421
7.8684 | -34.2415 | -13.552
7.8745 | -34.2419 | -13.5619
7.8806 | -34.2422 | -13.5719
7.8867 | -34.2426 | -13.5814
7.8928 | -34.243 | -13.591
7.8989 | -34.2433 | -13.6009
7.905 | -34.2437 | -13.6108
7.9111 | -34.244 | -13.6206
7.9172 | -34.2444 | -13.6305
7.9233 | -34.2447 | -13.6403
7.9294 | -34.2451 | -13.6503
7.9355 | -34.2454 | -13.66
7.9416 | -34.2457 | -13.6699
7.9477 | -34.2461 | -13.6796
7.9539 | -34.2464 | -13.6893
7.96 | -34.2468 | -13.6992
7.9661 | -34.2472 | -13.709
7.9722 | -34.2475 | -13.7186
7.9783 | -34.2479 | -13.7287
7.9844 | -34.2482 | -13.7384
7.9905 | -34.2486 | -13.7484
7.9966 | -34.2489 | -13.7584
8.0027 | -34.2493 | -13.7682
8.0088 | -34.2496 | -13.7781
8.0149 | -34.2499 | -13.7879
8.021 | -34.2503 | -13.7978
8.0271 | -34.2507 | -13.8078
8.0332 | -34.251 | -13.8177
8.0393 | -34.2514 | -13.8276
8.0454 | -34.2517 | -13.8376
8.0515 | -34.2521 | -13.8475
8.0576 | -34.2525 | -13.8577
8.0637 | -34.2528 | -13.8674
8.0698 | -34.2531 | -13.8771
8.0759 | -34.2534 | -13.8869
8.082 | -34.2538 | -13.8966
8.0881 | -34.2542 | -13.9065
8.0942 | -34.2545 | -13.9164
8.1004 | -34.2549 | -13.9263
8.1065 | -34.2553 | -13.936
8.1126 | -34.2556 | -13.946
8.1187 | -34.2559 | -13.956
8.1248 | -34.2563 | -13.9657
8.1309 | -34.2567 | -13.9754
8.137 | -34.257 | -13.985
8.1431 | -34.2574 | -13.9951
8.1492 | -34.2577 | -14.0049
8.1553 | -34.258 | -14.0148
8.1614 | -34.2584 | -14.0248
8.1675 | -34.2587 | -14.0346
8.1736 | -34.259 | -14.0447
8.1797 | -34.2594 | -14.0545
8.1858 | -34.2597 | -14.0644
8.1919 | -34.2601 | -14.0743
8.198 | -34.2604 | -14.0839
8.2041 | -34.2608 | -14.0939
8.2102 | -34.2611 | -14.1038
8.2163 | -34.2615 | -14.1135
8.2224 | -34.2618 | -14.1234
8.2285 | -34.2622 | -14.133
8.2346 | -34.2625 | -14.1429
8.2408 | -34.2629 | -14.1526
8.2469 | -34.2632 | -14.1625
8.253 | -34.2636 | -14.1722
8.2591 | -34.2639 | -14.1821
8.2652 | -34.2643 | -14.1921
8.2713 | -34.2646 | -14.202
8.2774 | -34.265 | -14.212
8.2835 | -34.2653 | -14.222
8.2896 | -34.2656 | -14.2316
8.2957 | -34.266 | -14.2416
8.3018 | -34.2663 | -14.2514
8.3079 | -34.2667 | -14.2612
8.314 | -34.2671 | -14.271
8.3201 | -34.2674 | -14.2808
8.3262 | -34.2678 | -14.2907
8.3323 | -34.2681 | -14.3005
8.3384 | -34.2685 | -14.3104
8.3445 | -34.2688 | -14.3202
8.3506 | -34.2692 | -14.3302
8.3567 | -34.2695 | -14.3401
8.3628 | -34.2698 | -14.35
8.3689 | -34.2702 | -14.3598
8.375 | -34.2705 | -14.3696
8.3812 | -34.2709 | -14.3793
8.3873 | -34.2712 | -14.3891
8.3934 | -34.2716 | -14.399
8.3995 | -34.2719 | -14.4087
8.4056 | -34.2722 | -14.4185
8.4117 | -34.2725 | -14.4284
8.4178 | -34.2729 | -14.4383
8.4239 | -34.2732 | -14.448
8.43 | -34.2736 | -14.4578
8.4361 | -34.2739 | -14.4677
8.4422 | -34.2742 | -14.4774
8.4483 | -34.2746 | -14.4872
8.4544 | -34.2749 | -14.4971
8.4605 | -34.2753 | -14.5067
8.4666 | -34.2756 | -14.5167
8.4727 | -34.2759 | -14.5266
8.4788 | -34.2763 | -14.5364
8.4849 | -34.2766 | -14.5462
8.491 | -34.277 | -14.556
8.4971 | -34.2773 | -14.5656
8.5032 | -34.2776 | -14.5754
8.5093 | -34.278 | -14.585
8.5154 | -34.2784 | -14.5949
8.5215 | -34.2787 | -14.6046
8.5277 | -34.279 | -14.6144
8.5338 | -34.2794 | -14.6242
8.5399 | -34.2797 | -14.634
8.546 | -34.2801 | -14.6438
8.5521 | -34.2804 | -14.6536
8.5582 | -34.2808 | -14.6634
8.5643 | -34.2811 | -14.6729
8.5704 | -34.2814 | -14.6825
8.5765 | -34.2818 | -14.6924
8.5826 | -34.2821 | -14.7018
8.5887 | -34.2825 | -14.7116
8.5948 | -34.2829 | -14.7214
8.6009 | -34.2832 | -14.731
8.607 | -34.2836 | -14.7409
8.6131 | -34.2839 | -14.7507
8.6192 | -34.2843 | -14.7604
8.6253 | -34.2846 | -14.77
8.6314 | -34.285 | -14.7797
8.6375 | -34.2853 | -14.7895
8.6436 | -34.2856 | -14.7992
8.6497 | -34.286 | -14.8089
8.6558 | -34.2863 | -14.8186
8.6619 | -34.2866 | -14.8284
8.6681 | -34.287 | -14.8384
8.6742 | -34.2873 | -14.848
8.6803 | -34.2877 | -14.8577
8.6864 | -34.288 | -14.8672
8.6925 | -34.2884 | -14.8767
8.6986 | -34.2887 | -14.8866
8.7047 | -34.2891 | -14.8964
8.7108 | -34.2894 | -14.9062
8.7169 | -34.2898 | -14.916
8.723 | -34.2901 | -14.9256
8.7291 | -34.2904 | -14.9352
8.7352 | -34.2908 | -14.9451
8.7413 | -34.2911 | -14.9549
8.7474 | -34.2915 | -14.9645
8.7535 | -34.2918 | -14.9743
8.7596 | -34.2922 | -14.9841
8.7657 | -34.2925 | -14.9939
8.7718 | -34.2928 | -15.0038
8.7779 | -34.2932 | -15.0136
8.784 | -34.2936 | -15.0234
8.7901 | -34.2939 | -15.0331
8.7962 | -34.2943 | -15.0426
8.8023 | -34.2946 | -15.0524
8.8084 | -34.295 | -15.062
8.8146 | -34.2953 | -15.0717
8.8207 | -34.2956 | -15.0815
8.8268 | -34.296 | -15.0909
8.8329 | -34.2963 | -15.1006
8.839 | -34.2967 | -15.1103
8.8451 | -34.297 | -15.1199
8.8512 | -34.2974 | -15.1296
8.8573 | -34.2977 | -15.1394
8.8634 | -34.2981 | -15.1489
8.8695 | -34.2985 | -15.1588
8.8756 | -34.2988 | -15.1686
8.8817 | -34.2991 | -15.1782
8.8878 | -34.2995 | -15.1881
8.8939 | -34.2998 | -15.1978
8.9 | -34.3002 | -15.2074
8.9061 | -34.3005 | -15.217
8.9122 | -34.3009 | -15.2267
8.9183 | -34.3013 | -15.2365
8.9244 | -34.3017 | -15.2464
8.9305 | -34.3021 | -15.2558
8.9366 | -34.3024 | -15.2654
8.9427 | -34.3028 | -15.2751
8.9488 | -34.3031 | -15.2848
8.955 | -34.3035 | -15.2944
8.9611 | -34.3038 | -15.3043
8.9672 | -34.3042 | -15.3139
8.9733 | -34.3046 | -15.3234
8.9794 | -34.305 | -15.333
8.9855 | -34.3053 | -15.3428
8.9916 | -34.3057 | -15.3526
8.9977 | -34.306 | -15.3621
9.0038 | -34.3065 | -15.3717
9.0099 | -34.3068 | -15.3814
9.016 | -34.3072 | -15.3911
9.0221 | -34.3075 | -15.4006
9.0282 | -34.3079 | -15.4101
9.0343 | -34.3083 | -15.4197
9.0404 | -34.3087 | -15.4295
9.0465 | -34.309 | -15.4393
9.0526 | -34.3094 | -15.4488
9.0587 | -34.3098 | -15.4584
9.0648 | -34.3102 | -15.4679
9.0709 | -34.3105 | -15.4775
9.077 | -34.3109 | -15.4872
9.0831 | -34.3113 | -15.4969
9.0892 | -34.3117 | -15.5067
9.0953 | -34.3121 | -15.5164
9.1015 | -34.3124 | -15.5261
9.1076 | -34.3128 | -15.5356
9.1137 | -34.3132 | -15.5452
9.1198 | -34.3136 | -15.5548
9.1259 | -34.3139 | -15.5645
9.132 | -34.3143 | -15.5743
9.1381 | -34.3147 | -15.5839
9.1442 | -34.3151 | -15.5933
9.1503 | -34.3155 | -15.603
9.1564 | -34.3159 | -15.6127
9.1625 | -34.3162 | -15.6224
9.1686 | -34.3166 | -15.632
9.1747 | -34.317 | -15.6416
9.1808 | -34.3174 | -15.6512
9.1869 | -34.3178 | -15.6609
9.193 | -34.3182 | -15.6706
9.1991 | -34.3185 | -15.6802
9.2052 | -34.3189 | -15.6898
9.2113 | -34.3193 | -15.6995
9.2174 | -34.3197 | -15.7091
9.2235 | -34.3201 | -15.7187
9.2296 | -34.3205 | -15.7283
9.2357 | -34.3209 | -15.7381
9.2419 | -34.3213 | -15.7477
9.248 | -34.3217 | -15.7573
9.2541 | -34.3221 | -15.7671
9.2602 | -34.3225 | -15.7768
9.2663 | -34.3229 | -15.7862
9.2724 | -34.3233 | -15.7957
9.2785 | -34.3236 | -15.8053
9.2846 | -34.3241 | -15.8148
9.2907 | -34.3245 | -15.8243
9.2968 | -34.3248 | -15.8341
9.3029 | -34.3252 | -15.8439
9.309 | -34.3256 | -15.8535
9.3151 | -34.326 | -15.863
9.3212 | -34.3264 | -15.8724
9.3273 | -34.3268 | -15.8819
9.3334 | -34.3272 | -15.8914
9.3395 | -34.3276 | -15.9009
9.3456 | -34.328 | -15.9103
9.3517 | -34.3284 | -15.9199
9.3578 | -34.3288 | -15.9296
9.3639 | -34.3292 | -15.939
9.37 | -34.3296 | -15.9486
9.3761 | -34.33 | -15.9584
9.3822 | -34.3304 | -15.9681
9.3884 | -34.3308 | -15.9775
9.3945 | -34.3312 | -15.9872
9.4006 | -34.3316 | -15.9969
9.4067 | -34.332 | -16.0065
9.4128 | -34.3324 | -16.0162
9.4189 | -34.3328 | -16.0259
9.425 | -34.3332 | -16.0354
9.4311 | -34.3336 | -16.045
9.4372 | -34.3341 | -16.0546
9.4433 | -34.3344 | -16.0642
9.4494 | -34.3348 | -16.0737
9.4555 | -34.3353 | -16.0834
9.4616 | -34.3357 | -16.0929
9.4677 | -34.3361 | -16.1025
9.4738 | -34.3365 | -16.112
9.4799 | -34.3369 | -16.1216
9.486 | -34.3373 | -16.1312
9.4921 | -34.3377 | -16.1407
9.4982 | -34.3381 | -16.15
9.5043 | -34.3385 | -16.1595
9.5104 | -34.3389 | -16.1691
9.5165 | -34.3393 | -16.1787
9.5226 | -34.3397 | -16.1885
9.5288 | -34.3401 | -16.198
9.5349 | -34.3405 | -16.2076
9.541 | -34.3409 | -16.2169
9.5471 | -34.3413 | -16.2264
9.5532 | -34.3417 | -16.2361
9.5593 | -34.3421 | -16.2458
9.5654 | -34.3425 | -16.2553
9.5715 | -34.3429 | -16.2648
9.5776 | -34.3433 | -16.2745
9.5837 | -34.3437 | -16.284
9.5898 | -34.3441 | -16.2937
9.5959 | -34.3445 | -16.3032
9.602 | -34.3449 | -16.3126
9.6081 | -34.3453 | -16.3221
9.6142 | -34.3458 | -16.3316
9.6203 | -34.3461 | -16.3413
9.6264 | -34.3465 | -16.3508
9.6325 | -34.3469 | -16.3605
9.6386 | -34.3473 | -16.3699
9.6447 | -34.3477 | -16.3795
9.6508 | -34.3482 | -16.389
9.6569 | -34.3486 | -16.3985
9.663 | -34.349 | -16.4081
9.6691 | -34.3494 | -16.4176
9.6753 | -34.3499 | -16.4271
9.6814 | -34.3503 | -16.4365
9.6875 | -34.3507 | -16.4458
9.6936 | -34.3511 | -16.4552
9.6997 | -34.3515 | -16.4649
9.7058 | -34.352 | -16.4743
9.7119 | -34.3524 | -16.484
9.718 | -34.3528 | -16.4937
9.7241 | -34.3532 | -16.5031
9.7302 | -34.3536 | -16.5127
9.7363 | -34.354 | -16.5221
9.7424 | -34.3545 | -16.5316
9.7485 | -34.3549 | -16.5411
9.7546 | -34.3553 | -16.5505
9.7607 | -34.3557 | -16.5603
9.7668 | -34.3561 | -16.5698
9.7729 | -34.3565 | -16.5792
9.779 | -34.3569 | -16.5886
9.7851 | -34.3574 | -16.5983
9.7912 | -34.3578 | -16.6077
9.7973 | -34.3582 | -16.6172
9.8034 | -34.3586 | -16.6264
9.8095 | -34.359 | -16.6361
9.8157 | -34.3595 | -16.6456
9.8218 | -34.3599 | -16.6554
9.8279 | -34.3603 | -16.6647
9.834 | -34.3607 | -16.6743
9.8401 | -34.3611 | -16.6838
9.8462 | -34.3616 | -16.693
9.8523 | -34.362 | -16.7024
9.8584 | -34.3624 | -16.7119
9.8645 | -34.3628 | -16.7215
9.8706 | -34.3632 | -16.7312
9.8767 | -34.3636 | -16.7406
9.8828 | -34.364 | -16.75
9.8889 | -34.3644 | -16.7593
9.895 | -34.3648 | -16.769
9.9011 | -34.3652 | -16.7784
9.9072 | -34.3656 | -16.7879
9.9133 | -34.366 | -16.7973
9.9194 | -34.3664 | -16.8069
9.9255 | -34.3668 | -16.8164
9.9316 | -34.3672 | -16.8258
9.9377 | -34.3677 | -16.8353
9.9438 | -34.3681 | -16.8449
9.9499 | -34.3685 | -16.8545
9.956 | -34.3689 | -16.864
9.9622 | -34.3693 | -16.8737
9.9683 | -34.3697 | -16.8832
9.9744 | -34.3701 | -16.8929
9.9805 | -34.3706 | -16.9024
9.9866 | -34.371 | -16.9119
9.9927 | -34.3714 | -16.9212
9.9988 | -34.3718 | -16.931
10.0049 | -34.3722 | -16.9403
10.011 | -34.3726 | -16.9498
10.0171 | -34.3731 | -16.9593
10.0232 | -34.3735 | -16.9686
10.0293 | -34.374 | -16.9781
10.0354 | -34.3744 | -16.9876
10.0415 | -34.3748 | -16.9972
10.0476 | -34.3752 | -17.0067
10.0537 | -34.3756 | -17.0163
10.0598 | -34.3761 | -17.0256
10.0659 | -34.3765 | -17.0351
10.072 | -34.3769 | -17.0444
10.0781 | -34.3773 | -17.0538
10.0842 | -34.3777 | -17.0634
10.0903 | -34.3781 | -17.0727
10.0964 | -34.3785 | -17.0819
10.1026 | -34.3789 | -17.0914
10.1087 | -34.3793 | -17.1008
10.1148 | -34.3797 | -17.1104
10.1209 | -34.3801 | -17.1197
10.127 | -34.3806 | -17.1291
10.1331 | -34.381 | -17.1384
10.1392 | -34.3814 | -17.1478
10.1453 | -34.3818 | -17.1571
10.1514 | -34.3822 | -17.1665
10.1575 | -34.3826 | -17.1761
10.1636 | -34.383 | -17.1853
10.1697 | -34.3835 | -17.1948
10.1758 | -34.3839 | -17.204
10.1819 | -34.3843 | -17.2134
10.188 | -34.3847 | -17.2229
10.1941 | -34.3851 | -17.2322
10.2002 | -34.3855 | -17.2416
10.2063 | -34.386 | -17.251
10.2124 | -34.3864 | -17.2606
10.2185 | -34.3868 | -17.2701
10.2246 | -34.3872 | -17.2795
10.2307 | -34.3876 | -17.2889
10.2368 | -34.3881 | -17.2981
10.2429 | -34.3885 | -17.3075
10.2491 | -34.3889 | -17.3171
10.2552 | -34.3894 | -17.3266
10.2613 | -34.3898 | -17.3361
10.2674 | -34.3902 | -17.3458
10.2735 | -34.3906 | -17.3554
10.2796 | -34.3911 | -17.3648
10.2857 | -34.3915 | -17.3743
10.2918 | -34.3919 | -17.3838
10.2979 | -34.3923 | -17.3933
10.304 | -34.3928 | -17.4024
10.3101 | -34.3932 | -17.4119
10.3162 | -34.3936 | -17.4213
10.3223 | -34.394 | -17.4306
10.3284 | -34.3945 | -17.4399
10.3345 | -34.3949 | -17.4492
10.3406 | -34.3953 | -17.4585
10.3467 | -34.3957 | -17.468
10.3528 | -34.3962 | -17.4773
10.3589 | -34.3966 | -17.4868
10.365 | -34.397 | -17.4965
10.3711 | -34.3974 | -17.5059
10.3772 | -34.3978 | -17.5152
10.3833 | -34.3983 | -17.5246
10.3895 | -34.3987 | -17.534
10.3956 | -34.3991 | -17.5435
10.4017 | -34.3996 | -17.5528
10.4078 | -34.3999 | -17.5622
10.4139 | -34.4004 | -17.5717
10.42 | -34.4008 | -17.5813
10.4261 | -34.4012 | -17.5907
10.4322 | -34.4017 | -17.6
10.4383 | -34.4021 | -17.6094
10.4444 | -34.4025 | -17.6187
10.4505 | -34.4029 | -17.6281
10.4566 | -34.4033 | -17.6373
10.4627 | -34.4037 | -17.6467
10.4688 | -34.4041 | -17.6559
10.4749 | -34.4045 | -17.6654
10.481 | -34.4049 | -17.6746
10.4871 | -34.4054 | -17.6837
10.4932 | -34.4058 | -17.6931
10.4993 | -34.4062 | -17.7024
10.5054 | -34.4066 | -17.7118
10.5115 | -34.407 | -17.7211
10.5176 | -34.4075 | -17.7302
10.5237 | -34.4079 | -17.7398
10.5298 | -34.4083 | -17.749
10.536 | -34.4087 | -17.7585
10.5421 | -34.4092 | -17.7677
10.5482 | -34.4096 | -17.7771
10.5543 | -34.41 | -17.7863
10.5604 | -34.4104 | -17.7954
10.5665 | -34.4108 | -17.8046
10.5726 | -34.4113 | -17.814
10.5787 | -34.4117 | -17.8233
10.5848 | -34.4121 | -17.8324
10.5909 | -34.4125 | -17.8418
10.597 | -34.4129 | -17.8512
10.6031 | -34.4133 | -17.8607
10.6092 | -34.4138 | -17.87
10.6153 | -34.4142 | -17.8794
10.6214 | -34.4146 | -17.8886
10.6275 | -34.415 | -17.898
10.6336 | -34.4154 | -17.9075
10.6397 | -34.4158 | -17.9168
10.6458 | -34.4163 | -17.9261
10.6519 | -34.4167 | -17.9355
10.658 | -34.4171 | -17.9447
10.6641 | -34.4175 | -17.9542
10.6702 | -34.4179 | -17.9635
10.6764 | -34.4184 | -17.9728
10.6825 | -34.4188 | -17.982
10.6886 | -34.4192 | -17.9912
10.6947 | -34.4197 | -18.0008
10.7008 | -34.4201 | -18.0102
10.7069 | -34.4205 | -18.0196
10.713 | -34.4209 | -18.0289
10.7191 | -34.4214 | -18.0382
10.7252 | -34.4218 | -18.0472
10.7313 | -34.4222 | -18.0564
10.7374 | -34.4226 | -18.0658
10.7435 | -34.4231 | -18.075
10.7496 | -34.4235 | -18.0841
10.7557 | -34.4239 | -18.0935
10.7618 | -34.4243 | -18.1026
10.7679 | -34.4248 | -18.112
10.774 | -34.4252 | -18.1214
10.7801 | -34.4256 | -18.1306
10.7862 | -34.426 | -18.1398
10.7923 | -34.4264 | -18.1492
10.7984 | -34.4268 | -18.1584
10.8045 | -34.4273 | -18.1677
10.8106 | -34.4277 | -18.1771
10.8168 | -34.4281 | -18.1862
10.8229 | -34.4285 | -18.1956
10.829 | -34.4289 | -18.2048
10.8351 | -34.4293 | -18.2138
10.8412 | -34.4298 | -18.223
10.8473 | -34.4302 | -18.2321
10.8534 | -34.4307 | -18.2413
10.8595 | -34.4311 | -18.2506
10.8656 | -34.4315 | -18.2597
10.8717 | -34.4319 | -18.2689
10.8778 | -34.4324 | -18.2783
10.8839 | -34.4328 | -18.2876
10.89 | -34.4332 | -18.2968
10.8961 | -34.4337 | -18.306
10.9022 | -34.4341 | -18.3152
10.9083 | -34.4346 | -18.3244
10.9144 | -34.435 | -18.3335
10.9205 | -34.4354 | -18.3427
10.9266 | -34.4358 | -18.3517
10.9327 | -34.4363 | -18.3608
10.9388 | -34.4367 | -18.3702
10.9449 | -34.4371 | -18.3794
10.951 | -34.4376 | -18.3885
10.9571 | -34.438 | -18.3978
10.9633 | -34.4384 | -18.407
10.9694 | -34.4388 | -18.4164
10.9755 | -34.4393 | -18.4255
10.9816 | -34.4397 | -18.4348
10.9877 | -34.4401 | -18.4439
10.9938 | -34.4406 | -18.4533
10.9999 | -34.441 | -18.4626
11.006 | -34.4414 | -18.4717
11.0121 | -34.4418 | -18.4809
11.0182 | -34.4423 | -18.4903
11.0243 | -34.4427 | -18.4995
11.0304 | -34.4431 | -18.5088
11.0365 | -34.4435 | -18.5179
11.0426 | -34.444 | -18.5271
11.0487 | -34.4444 | -18.5363
11.0548 | -34.4448 | -18.5453
11.0609 | -34.4453 | -18.5546
11.067 | -34.4457 | -18.5638
11.0731 | -34.4462 | -18.5728
11.0792 | -34.4466 | -18.582
11.0853 | -34.447 | -18.5911
11.0914 | -34.4475 | -18.6005
11.0975 | -34.4479 | -18.6095
11.1037 | -34.4483 | -18.6187
11.1098 | -34.4487 | -18.6279
11.1159 | -34.4492 | -18.6368
11.122 | -34.4496 | -18.6459
11.1281 | -34.45 | -18.6553
11.1342 | -34.4505 | -18.6645
11.1403 | -34.4509 | -18.6737
11.1464 | -34.4513 | -18.6829
11.1525 | -34.4518 | -18.6919
11.1586 | -34.4522 | -18.7011
11.1647 | -34.4526 | -18.7103
11.1708 | -34.453 | -18.7196
11.1769 | -34.4534 | -18.7288
11.183 | -34.4539 | -18.7378
11.1891 | -34.4543 | -18.7469
11.1952 | -34.4548 | -18.7559
11.2013 | -34.4552 | -18.765
11.2074 | -34.4556 | -18.774
11.2135 | -34.456 | -18.7833
11.2196 | -34.4565 | -18.7923
11.2257 | -34.4569 | -18.8015
11.2318 | -34.4574 | -18.8107
11.2379 | -34.4578 | -18.8199
11.244 | -34.4583 | -18.8291
11.2502 | -34.4587 | -18.8383
11.2563 | -34.4592 | -18.8475
11.2624 | -34.4596 | -18.8564
11.2685 | -34.46 | -18.8656
11.2746 | -34.4605 | -18.8747
11.2807 | -34.4609 | -18.884
11.2868 | -34.4614 | -18.8933
11.2929 | -34.4618 | -18.9022
11.299 | -34.4623 | -18.9115
11.3051 | -34.4627 | -18.9209
11.3112 | -34.4631 | -18.93
11.3173 | -34.4636 | -18.9391
11.3234 | -34.464 | -18.948
11.3295 | -34.4644 | -18.9574
11.3356 | -34.4649 | -18.9669
11.3417 | -34.4653 | -18.9759
11.3478 | -34.4657 | -18.9851
11.3539 | -34.4662 | -18.9942
11.36 | -34.4666 | -19.0032
11.3661 | -34.4671 | -19.0122
11.3722 | -34.4675 | -19.0211
11.3783 | -34.4679 | -19.0304
11.3844 | -34.4684 | -19.0395
11.3906 | -34.4688 | -19.0486
11.3967 | -34.4693 | -19.0575
11.4028 | -34.4697 | -19.0667
11.4089 | -34.4701 | -19.0756
11.415 | -34.4705 | -19.0849
11.4211 | -34.4709 | -19.0942
11.4272 | -34.4714 | -19.1033
11.4333 | -34.4718 | -19.1123
11.4394 | -34.4722 | -19.1214
11.4455 | -34.4727 | -19.1304
11.4516 | -34.4731 | -19.1395
11.4577 | -34.4735 | -19.1486
11.4638 | -34.474 | -19.1576
11.4699 | -34.4744 | -19.1668
11.476 | -34.4749 | -19.1761
11.4821 | -34.4753 | -19.1851
11.4882 | -34.4757 | -19.1944
11.4943 | -34.4762 | -19.2034
11.5004 | -34.4766 | -19.2123
11.5065 | -34.4771 | -19.2213
11.5126 | -34.4775 | -19.2304
11.5187 | -34.4779 | -19.2394
11.5248 | -34.4783 | -19.2487
11.5309 | -34.4788 | -19.2578
11.5371 | -34.4792 | -19.2669
11.5432 | -34.4796 | -19.2759
11.5493 | -34.48 | -19.2851
11.5554 | -34.4805 | -19.2942
11.5615 | -34.4809 | -19.3033
11.5676 | -34.4813 | -19.3124
11.5737 | -34.4818 | -19.3214
11.5798 | -34.4822 | -19.3303
11.5859 | -34.4827 | -19.3394
11.592 | -34.4831 | -19.3484
11.5981 | -34.4835 | -19.3574
11.6042 | -34.4839 | -19.3665
11.6103 | -34.4843 | -19.3754
11.6164 | -34.4848 | -19.3844
11.6225 | -34.4852 | -19.3934
11.6286 | -34.4857 | -19.4024
11.6347 | -34.4861 | -19.4115
11.6408 | -34.4865 | -19.4208
11.6469 | -34.487 | -19.4299
11.653 | -34.4874 | -19.439
11.6591 | -34.4878 | -19.4479
11.6652 | -34.4883 | -19.4569
11.6713 | -34.4887 | -19.4657
11.6775 | -34.4892 | -19.4748
11.6836 | -34.4896 | -19.4837
11.6897 | -34.4901 | -19.4928
11.6958 | -34.4905 | -19.5019
11.7019 | -34.491 | -19.511
11.708 | -34.4914 | -19.5201
11.7141 | -34.4919 | -19.5289
11.7202 | -34.4923 | -19.538
11.7263 | -34.4927 | -19.5472
11.7324 | -34.4931 | -19.5565
11.7385 | -34.4935 | -19.5656
11.7446 | -34.4939 | -19.5744
11.7507 | -34.4944 | -19.5831
11.7568 | -34.4948 | -19.5922
11.7629 | -34.4953 | -19.6013
11.769 | -34.4956 | -19.6102
11.7751 | -34.4961 | -19.6192
11.7812 | -34.4965 | -19.628
11.7873 | -34.4969 | -19.6371
11.7934 | -34.4974 | -19.6463
11.7995 | -34.4978 | -19.6553
11.8056 | -34.4982 | -19.6643
11.8117 | -34.4987 | -19.6732
11.8178 | -34.4991 | -19.6821
11.824 | -34.4995 | -19.691
11.8301 | -34.4999 | -19.6999
11.8362 | -34.5004 | -19.7087
11.8423 | -34.5008 | -19.7176
11.8484 | -34.5012 | -19.7266
11.8545 | -34.5016 | -19.7354
11.8606 | -34.5021 | -19.7442
11.8667 | -34.5025 | -19.7532
11.8728 | -34.5029 | -19.7619
11.8789 | -34.5034 | -19.7708
11.885 | -34.5038 | -19.7798
11.8911 | -34.5042 | -19.7887
11.8972 | -34.5047 | -19.7976
11.9033 | -34.5051 | -19.8064
11.9094 | -34.5055 | -19.8154
11.9155 | -34.5059 | -19.8244
11.9216 | -34.5064 | -19.8332
11.9277 | -34.5068 | -19.8419
11.9338 | -34.5073 | -19.8508
11.9399 | -34.5077 | -19.8599
11.946 | -34.5081 | -19.8687
11.9521 | -34.5086 | -19.8776
11.9582 | -34.509 | -19.8864
11.9644 | -34.5094 | -19.8953
11.9705 | -34.5099 | -19.9043
11.9766 | -34.5103 | -19.9132
11.9827 | -34.5108 | -19.9221
11.9888 | -34.5112 | -19.9309
11.9949 | -34.5117 | -19.9394
12.001 | -34.5121 | -19.9483
12.0071 | -34.5125 | -19.9571
12.0132 | -34.5129 | -19.9658
12.0193 | -34.5134 | -19.9747
12.0254 | -34.5137 | -19.9836
12.0315 | -34.5142 | -19.9923
12.0376 | -34.5146 | -20.0011
12.0437 | -34.515 | -20.0099
12.0498 | -34.5155 | -20.0188
12.0559 | -34.5159 | -20.0277
12.062 | -34.5164 | -20.0363
12.0681 | -34.5168 | -20.0452
12.0742 | -34.5173 | -20.0539
12.0803 | -34.5177 | -20.0628
12.0864 | -34.5181 | -20.0717
12.0925 | -34.5185 | -20.0808
12.0986 | -34.519 | -20.0896
12.1047 | -34.5194 | -20.0983
12.1109 | -34.5199 | -20.1072
12.117 | -34.5203 | -20.116
12.1231 | -34.5208 | -20.1248
12.1292 | -34.5212 | -20.1334
12.1353 | -34.5216 | -20.142
12.1414 | -34.5221 | -20.1508
12.1475 | -34.5225 | -20.1597
12.1536 | -34.5229 | -20.1687
12.1597 | -34.5233 | -20.1775
12.1658 | -34.5238 | -20.1862
12.1719 | -34.5242 | -20.1949
12.178 | -34.5246 | -20.2037
12.1841 | -34.5251 | -20.2126
12.1902 | -34.5255 | -20.2214
12.1963 | -34.526 | -20.2303
12.2024 | -34.5264 | -20.2389
12.2085 | -34.5268 | -20.2477
12.2146 | -34.5273 | -20.2564
12.2207 | -34.5277 | -20.2651
12.2268 | -34.5282 | -20.2739
12.2329 | -34.5286 | -20.2827
12.239 | -34.529 | -20.2911
12.2451 | -34.5295 | -20.3
12.2513 | -34.5299 | -20.3089
12.2574 | -34.5303 | -20.3177
12.2635 | -34.5308 | -20.3264
12.2696 | -34.5312 | -20.3352
12.2757 | -34.5317 | -20.3435
12.2818 | -34.5321 | -20.3524
12.2879 | -34.5325 | -20.3611
12.294 | -34.5329 | -20.3698
12.3001 | -34.5334 | -20.3786
12.3062 | -34.5339 | -20.3872
12.3123 | -34.5343 | -20.3956
12.3184 | -34.5347 | -20.4042
12.3245 | -34.5352 | -20.4128
12.3306 | -34.5356 | -20.4217
12.3367 | -34.5361 | -20.4304
12.3428 | -34.5365 | -20.4391
12.3489 | -34.537 | -20.4475
12.355 | -34.5374 | -20.4564
12.3611 | -34.5378 | -20.4652
12.3672 | -34.5383 | -20.4739
12.3733 | -34.5387 | -20.4826
12.3794 | -34.5391 | -20.4911
12.3855 | -34.5396 | -20.4996
12.3916 | -34.54 | -20.5083
12.3978 | -34.5405 | -20.5169
12.4039 | -34.5409 | -20.5256
12.41 | -34.5414 | -20.5343
12.4161 | -34.5418 | -20.5425
12.4222 | -34.5423 | -20.5513
12.4283 | -34.5427 | -20.5599
12.4344 | -34.5431 | -20.5686
12.4405 | -34.5436 | -20.5772
12.4466 | -34.544 | -20.5859
12.4527 | -34.5445 | -20.5943
12.4588 | -34.5449 | -20.603
12.4649 | -34.5454 | -20.6117
12.471 | -34.5458 | -20.6205
12.4771 | -34.5462 | -20.629
12.4832 | -34.5467 | -20.6376
12.4893 | -34.5471 | -20.6459
12.4954 | -34.5476 | -20.6544
12.5015 | -34.548 | -20.6629
12.5076 | -34.5484 | -20.6717
12.5137 | -34.5489 | -20.6802
12.5198 | -34.5494 | -20.6886
12.5259 | -34.5498 | -20.6972
12.532 | -34.5502 | -20.706
12.5382 | -34.5507 | -20.7147
12.5443 | -34.5511 | -20.7233
12.5504 | -34.5515 | -20.732
12.5565 | -34.552 | -20.7406
12.5626 | -34.5524 | -20.7493
12.5687 | -34.5529 | -20.7579
12.5748 | -34.5533 | -20.7665
12.5809 | -34.5537 | -20.775
12.587 | -34.5542 | -20.7837
12.5931 | -34.5547 | -20.7921
12.5992 | -34.5551 | -20.8007
12.6053 | -34.5555 | -20.8092
12.6114 | -34.556 | -20.8175
12.6175 | -34.5564 | -20.8262
12.6236 | -34.5568 | -20.8347
12.6297 | -34.5573 | -20.8433
12.6358 | -34.5577 | -20.8517
12.6419 | -34.5582 | -20.8603
12.648 | -34.5586 | -20.8689
12.6541 | -34.5591 | -20.8776
12.6602 | -34.5595 | -20.886
12.6663 | -34.5599 | -20.8947
12.6724 | -34.5604 | -20.9031
12.6785 | -34.5609 | -20.9116
12.6847 | -34.5613 | -20.9201
12.6908 | -34.5618 | -20.9287
12.6969 | -34.5623 | -20.9371
12.703 | -34.5627 | -20.9458
12.7091 | -34.5631 | -20.9544
12.7152 | -34.5636 | -20.9628
12.7213 | -34.564 | -20.9713
12.7274 | -34.5645 | -20.9799
12.7335 | -34.5649 | -20.9884
12.7396 | -34.5653 | -20.9968
12.7457 | -34.5658 | -21.0054
12.7518 | -34.5662 | -21.0139
12.7579 | -34.5667 | -21.0226
12.764 | -34.5671 | -21.0311
12.7701 | -34.5676 | -21.0395
12.7762 | -34.568 | -21.0481
12.7823 | -34.5685 | -21.0566
12.7884 | -34.569 | -21.0653
12.7945 | -34.5694 | -21.074
12.8006 | -34.5699 | -21.0825
12.8067 | -34.5703 | -21.091
12.8128 | -34.5708 | -21.0994
12.8189 | -34.5712 | -21.1079
12.8251 | -34.5716 | -21.1164
12.8312 | -34.5721 | -21.125
12.8373 | -34.5725 | -21.1332
12.8434 | -34.573 | -21.1417
12.8495 | -34.5734 | -21.1501
12.8556 | -34.5739 | -21.1584
12.8617 | -34.5743 | -21.1669
12.8678 | -34.5747 | -21.1757
12.8739 | -34.5752 | -21.1841
12.88 | -34.5756 | -21.1925
12.8861 | -34.576 | -21.201
12.8922 | -34.5765 | -21.2096
12.8983 | -34.577 | -21.2181
12.9044 | -34.5774 | -21.2265
12.9105 | -34.5779 | -21.2348
12.9166 | -34.5783 | -21.2433
12.9227 | -34.5788 | -21.2517
12.9288 | -34.5792 | -21.2603
12.9349 | -34.5797 | -21.2687
12.941 | -34.5801 | -21.2771
12.9471 | -34.5806 | -21.2856
12.9532 | -34.5811 | -21.2941
12.9593 | -34.5815 | -21.3028
12.9654 | -34.582 | -21.3114
12.9716 | -34.5824 | -21.3199
12.9777 | -34.5829 | -21.3282
12.9838 | -34.5834 | -21.3366
12.9899 | -34.5838 | -21.345
12.996 | -34.5842 | -21.3535
13.0021 | -34.5847 | -21.3619
13.0082 | -34.5851 | -21.3702
13.0143 | -34.5856 | -21.3784
13.0204 | -34.5861 | -21.3867
13.0265 | -34.5865 | -21.395
13.0326 | -34.587 | -21.4035
13.0387 | -34.5874 | -21.4122
13.0448 | -34.5879 | -21.4206
13.0509 | -34.5884 | -21.4291
13.057 | -34.5888 | -21.4374
13.0631 | -34.5893 | -21.4458
13.0692 | -34.5897 | -21.4541
13.0753 | -34.5902 | -21.4626
13.0814 | -34.5907 | -21.4711
13.0875 | -34.5912 | -21.4793
13.0936 | -34.5916 | -21.4878
13.0997 | -34.5921 | -21.4961
13.1058 | -34.5926 | -21.5045
13.112 | -34.593 | -21.5129
13.1181 | -34.5935 | -21.5212
13.1242 | -34.5939 | -21.5296
13.1303 | -34.5944 | -21.5379
13.1364 | -34.5948 | -21.5465
13.1425 | -34.5953 | -21.5548
13.1486 | -34.5957 | -21.5633
13.1547 | -34.5962 | -21.5716
13.1608 | -34.5966 | -21.5801
13.1669 | -34.5971 | -21.5883
13.173 | -34.5975 | -21.5968
13.1791 | -34.598 | -21.6051
13.1852 | -34.5984 | -21.6136
13.1913 | -34.5989 | -21.6218
13.1974 | -34.5993 | -21.6301
13.2035 | -34.5998 | -21.6386
13.2096 | -34.6003 | -21.6468
13.2157 | -34.6008 | -21.6553
13.2218 | -34.6013 | -21.6638
13.2279 | -34.6017 | -21.6719
13.234 | -34.6021 | -21.6802
13.2401 | -34.6026 | -21.6886
13.2462 | -34.6031 | -21.6968
13.2524 | -34.6035 | -21.7053
13.2585 | -34.604 | -21.7139
13.2646 | -34.6044 | -21.7223
13.2707 | -34.6049 | -21.7307
13.2768 | -34.6054 | -21.7391
13.2829 | -34.6058 | -21.7476
13.289 | -34.6063 | -21.7559
13.2951 | -34.6067 | -21.7643
13.3012 | -34.6072 | -21.7726
13.3073 | -34.6077 | -21.7811
13.3134 | -34.6081 | -21.7895
13.3195 | -34.6086 | -21.798
13.3256 | -34.609 | -21.806
13.3317 | -34.6095 | -21.8147
13.3378 | -34.6099 | -21.823
13.3439 | -34.6104 | -21.8315
13.35 | -34.6108 | -21.8399
13.3561 | -34.6113 | -21.848
13.3622 | -34.6117 | -21.8561
13.3683 | -34.6122 | -21.8641
13.3744 | -34.6127 | -21.8721
13.3805 | -34.6131 | -21.8804
13.3866 | -34.6136 | -21.8888
13.3927 | -34.614 | -21.897
13.3989 | -34.6145 | -21.9053
13.405 | -34.615 | -21.9136
13.4111 | -34.6154 | -21.9217
13.4172 | -34.6158 | -21.9301
13.4233 | -34.6162 | -21.9384
13.4294 | -34.6167 | -21.9469
13.4355 | -34.6171 | -21.9551
13.4416 | -34.6176 | -21.9636
13.4477 | -34.618 | -21.9719
13.4538 | -34.6185 | -21.9802
13.4599 | -34.619 | -21.9886
13.466 | -34.6194 | -21.9971
13.4721 | -34.6199 | -22.0052
13.4782 | -34.6203 | -22.0135
13.4843 | -34.6208 | -22.0219
13.4904 | -34.6213 | -22.0299
13.4965 | -34.6218 | -22.0386
13.5026 | -34.6222 | -22.0469
13.5087 | -34.6227 | -22.0549
13.5148 | -34.6231 | -22.0634
13.5209 | -34.6236 | -22.0718
13.527 | -34.624 | -22.0799
13.5331 | -34.6245 | -22.088
13.5393 | -34.625 | -22.0961
13.5454 | -34.6254 | -22.1044
13.5515 | -34.6259 | -22.1128
13.5576 | -34.6263 | -22.121
13.5637 | -34.6268 | -22.1294
13.5698 | -34.6273 | -22.1378
13.5759 | -34.6277 | -22.1457
13.582 | -34.6282 | -22.1541
13.5881 | -34.6287 | -22.1624
13.5942 | -34.6292 | -22.1708
13.6003 | -34.6297 | -22.179
13.6064 | -34.6301 | -22.1873
13.6125 | -34.6306 | -22.1959
13.6186 | -34.631 | -22.2039
13.6247 | -34.6315 | -22.2124
13.6308 | -34.632 | -22.2208
13.6369 | -34.6324 | -22.2293
13.643 | -34.6329 | -22.2376
13.6491 | -34.6334 | -22.2458
13.6552 | -34.6339 | -22.254
13.6613 | -34.6343 | -22.2623
13.6674 | -34.6348 | -22.2706
13.6735 | -34.6353 | -22.2789
13.6796 | -34.6357 | -22.2872
13.6858 | -34.6362 | -22.2955
13.6919 | -34.6367 | -22.3038
13.698 | -34.6371 | -22.3121
13.7041 | -34.6376 | -22.3205
13.7102 | -34.6381 | -22.3287
13.7163 | -34.6386 | -22.3369
13.7224 | -34.639 | -22.3453
13.7285 | -34.6395 | -22.3536
13.7346 | -34.6399 | -22.3619
13.7407 | -34.6404 | -22.3701
13.7468 | -34.6408 | -22.3784
13.7529 | -34.6413 | -22.3866
13.759 | -34.6417 | -22.3949
13.7651 | -34.6422 | -22.4032
13.7712 | -34.6427 | -22.4115
13.7773 | -34.6431 | -22.4197
13.7834 | -34.6436 | -22.4279
13.7895 | -34.6441 | -22.4362
13.7956 | -34.6446 | -22.4446
13.8017 | -34.645 | -22.4529
13.8078 | -34.6455 | -22.4611
13.8139 | -34.6459 | -22.4693
13.82 | -34.6464 | -22.4774
13.8262 | -34.6468 | -22.4857
13.8323 | -34.6473 | -22.494
13.8384 | -34.6478 | -22.5024
13.8445 | -34.6483 | -22.5106
13.8506 | -34.6487 | -22.519
13.8567 | -34.6492 | -22.5271
13.8628 | -34.6497 | -22.5355
13.8689 | -34.6502 | -22.5439
13.875 | -34.6506 | -22.5521
13.8811 | -34.6511 | -22.5605
13.8872 | -34.6515 | -22.5689
13.8933 | -34.652 | -22.5771
13.8994 | -34.6524 | -22.5854
13.9055 | -34.6529 | -22.5938
13.9116 | -34.6534 | -22.6021
13.9177 | -34.6538 | -22.6105
13.9238 | -34.6543 | -22.6188
13.9299 | -34.6548 | -22.627
13.936 | -34.6552 | -22.6353
13.9421 | -34.6557 | -22.6436
13.9482 | -34.6561 | -22.6518
13.9543 | -34.6566 | -22.6603
13.9604 | -34.657 | -22.6684
13.9665 | -34.6575 | -22.6765
13.9727 | -34.658 | -22.6849
13.9788 | -34.6585 | -22.693
13.9849 | -34.6589 | -22.7014
13.991 | -34.6594 | -22.7097
13.9971 | -34.6599 | -22.7178
14.0032 | -34.6603 | -22.7261
14.0093 | -34.6608 | -22.7344
14.0154 | -34.6613 | -22.7426
14.0215 | -34.6617 | -22.7508
14.0276 | -34.6622 | -22.759
14.0337 | -34.6627 | -22.7672
14.0398 | -34.6632 | -22.7755
14.0459 | -34.6636 | -22.7838
14.052 | -34.6641 | -22.7921
14.0581 | -34.6645 | -22.8004
14.0642 | -34.665 | -22.8087
14.0703 | -34.6655 | -22.817
14.0764 | -34.666 | -22.8253
14.0825 | -34.6664 | -22.8335
14.0886 | -34.6669 | -22.8418
14.0947 | -34.6674 | -22.8501
14.1008 | -34.6679 | -22.8582
14.1069 | -34.6684 | -22.8663
14.1131 | -34.6688 | -22.8747
14.1192 | -34.6693 | -22.8828
14.1253 | -34.6697 | -22.891
14.1314 | -34.6702 | -22.8993
14.1375 | -34.6707 | -22.9076
14.1436 | -34.6712 | -22.9158
14.1497 | -34.6717 | -22.9239
14.1558 | -34.6721 | -22.9322
14.1619 | -34.6726 | -22.9405
14.168 | -34.673 | -22.9488
14.1741 | -34.6735 | -22.9569
14.1802 | -34.674 | -22.965
14.1863 | -34.6745 | -22.9733
14.1924 | -34.6749 | -22.9815
14.1985 | -34.6754 | -22.9898
14.2046 | -34.6759 | -22.9981
14.2107 | -34.6764 | -23.0061
14.2168 | -34.6768 | -23.0143
14.2229 | -34.6773 | -23.0227
14.229 | -34.6778 | -23.0312
14.2351 | -34.6782 | -23.0393
14.2412 | -34.6787 | -23.0475
14.2473 | -34.6792 | -23.0555
14.2534 | -34.6797 | -23.0637
14.2596 | -34.6801 | -23.0719
14.2657 | -34.6805 | -23.08
14.2718 | -34.681 | -23.0882
14.2779 | -34.6815 | -23.0965
14.284 | -34.682 | -23.1048
14.2901 | -34.6825 | -23.1131
14.2962 | -34.683 | -23.1214
14.3023 | -34.6834 | -23.1297
14.3084 | -34.6839 | -23.1379
14.3145 | -34.6843 | -23.1461
14.3206 | -34.6848 | -23.1547
14.3267 | -34.6853 | -23.163
14.3328 | -34.6857 | -23.1711
14.3389 | -34.6862 | -23.1794
14.345 | -34.6866 | -23.1875
14.3511 | -34.6871 | -23.1959
14.3572 | -34.6876 | -23.2043
14.3633 | -34.6881 | -23.2127
14.3694 | -34.6885 | -23.2207
14.3755 | -34.6891 | -23.229
14.3816 | -34.6895 | -23.2373
14.3877 | -34.69 | -23.2454
14.3938 | -34.6905 | -23.2538
14.4 | -34.691 | -23.2618
14.4061 | -34.6914 | -23.2701
14.4122 | -34.6919 | -23.2783
14.4183 | -34.6923 | -23.2865
14.4244 | -34.6928 | -23.2948
14.4305 | -34.6933 | -23.3031
14.4366 | -34.6938 | -23.3113
14.4427 | -34.6943 | -23.3196
14.4488 | -34.6948 | -23.3279
14.4549 | -34.6952 | -23.336
14.461 | -34.6957 | -23.3442
14.4671 | -34.6962 | -23.3526
14.4732 | -34.6967 | -23.3609
14.4793 | -34.6971 | -23.3692
14.4854 | -34.6976 | -23.3774
14.4915 | -34.6981 | -23.3854
14.4976 | -34.6986 | -23.3937
14.5037 | -34.699 | -23.4018
14.5098 | -34.6995 | -23.4099
14.5159 | -34.7 | -23.4181
14.522 | -34.7004 | -23.4262
14.5281 | -34.7009 | -23.4343
14.5342 | -34.7014 | -23.4427
14.5403 | -34.7019 | -23.4508
14.5465 | -34.7023 | -23.459
14.5526 | -34.7028 | -23.4675
14.5587 | -34.7032 | -23.4757
14.5648 | -34.7037 | -23.4839
14.5709 | -34.7042 | -23.492
14.577 | -34.7046 | -23.5001
14.5831 | -34.7051 | -23.5083
14.5892 | -34.7056 | -23.5163
14.5953 | -34.706 | -23.5247
14.6014 | -34.7065 | -23.5328
14.6075 | -34.707 | -23.5409
14.6136 | -34.7074 | -23.549
14.6197 | -34.7079 | -23.5572
14.6258 | -34.7084 | -23.5655
14.6319 | -34.7088 | -23.5738
14.638 | -34.7093 | -23.582
14.6441 | -34.7098 | -23.5904
14.6502 | -34.7102 | -23.5988
14.6563 | -34.7107 | -23.6069
14.6624 | -34.7112 | -23.6152
14.6685 | -34.7117 | -23.6234
14.6746 | -34.7122 | -23.6317
14.6807 | -34.7126 | -23.64
14.6869 | -34.7131 | -23.6481
14.693 | -34.7136 | -23.6563
14.6991 | -34.714 | -23.6645
14.7052 | -34.7145 | -23.6725
14.7113 | -34.715 | -23.6808
14.7174 | -34.7155 | -23.6891
14.7235 | -34.716 | -23.6971
14.7296 | -34.7164 | -23.7051
14.7357 | -34.7169 | -23.7131
14.7418 | -34.7174 | -23.7213
14.7479 | -34.7179 | -23.7294
14.754 | -34.7184 | -23.7376
14.7601 | -34.7189 | -23.7457
14.7662 | -34.7193 | -23.7537
14.7723 | -34.7198 | -23.7617
14.7784 | -34.7203 | -23.7699
14.7845 | -34.7208 | -23.7779
14.7906 | -34.7213 | -23.7862
14.7967 | -34.7217 | -23.7944
14.8028 | -34.7222 | -23.8023
14.8089 | -34.7227 | -23.8104
14.815 | -34.7232 | -23.8184
14.8211 | -34.7236 | -23.8265
14.8272 | -34.7241 | -23.835
14.8334 | -34.7246 | -23.8431
14.8395 | -34.7251 | -23.8511
14.8456 | -34.7256 | -23.8594
14.8517 | -34.7261 | -23.8676
14.8578 | -34.7265 | -23.8759
14.8639 | -34.727 | -23.884
14.87 | -34.7275 | -23.8923
14.8761 | -34.728 | -23.9002
14.8822 | -34.7285 | -23.9084
14.8883 | -34.729 | -23.9167
14.8944 | -34.7294 | -23.9249
14.9005 | -34.7299 | -23.9332
14.9066 | -34.7304 | -23.9415
14.9127 | -34.7309 | -23.9493
14.9188 | -34.7314 | -23.9574
14.9249 | -34.7318 | -23.9657
14.931 | -34.7323 | -23.9738
14.9371 | -34.7328 | -23.9819
14.9432 | -34.7333 | -23.9901
14.9493 | -34.7338 | -23.9984
14.9554 | -34.7343 | -24.0066
14.9615 | -34.7348 | -24.0147
14.9676 | -34.7352 | -24.023
14.9738 | -34.7357 | -24.0311
14.9799 | -34.7362 | -24.0393
14.986 | -34.7367 | -24.0474
14.9921 | -34.7372 | -24.0556
14.9982 | -34.7377 | -24.0636
15.0043 | -34.7382 | -24.0718
15.0104 | -34.7386 | -24.0801
15.0165 | -34.7391 | -24.0886
15.0226 | -34.7396 | -24.0966
15.0287 | -34.7401 | -24.1047
15.0348 | -34.7405 | -24.1128
15.0409 | -34.741 | -24.1211
15.047 | -34.7415 | -24.1294
15.0531 | -34.742 | -24.1376
15.0592 | -34.7425 | -24.1458
15.0653 | -34.743 | -24.154
15.0714 | -34.7435 | -24.162
15.0775 | -34.744 | -24.1701
15.0836 | -34.7445 | -24.1782
15.0897 | -34.745 | -24.1862
15.0958 | -34.7454 | -24.1942
15.1019 | -34.7459 | -24.2025
15.108 | -34.7465 | -24.2105
15.1141 | -34.7469 | -24.2186
15.1203 | -34.7474 | -24.2268
15.1264 | -34.7479 | -24.2349
15.1325 | -34.7484 | -24.2429
15.1386 | -34.7489 | -24.2511
15.1447 | -34.7493 | -24.2593
15.1508 | -34.7498 | -24.2675
15.1569 | -34.7503 | -24.2756
15.163 | -34.7508 | -24.2835
15.1691 | -34.7514 | -24.2917
15.1752 | -34.7518 | -24.2999
15.1813 | -34.7523 | -24.308
15.1874 | -34.7528 | -24.3163
15.1935 | -34.7532 | -24.3242
15.1996 | -34.7537 | -24.3322
15.2057 | -34.7542 | -24.3401
15.2118 | -34.7547 | -24.3482
15.2179 | -34.7552 | -24.3563
15.224 | -34.7557 | -24.3642
15.2301 | -34.7562 | -24.3721
15.2362 | -34.7567 | -24.3804
15.2423 | -34.7572 | -24.3886
15.2484 | -34.7577 | -24.3965
15.2545 | -34.7581 | -24.4047
15.2607 | -34.7586 | -24.4129
15.2668 | -34.7592 | -24.421
15.2729 | -34.7596 | -24.4289
15.279 | -34.7601 | -24.4372
15.2851 | -34.7606 | -24.4453
15.2912 | -34.7611 | -24.4533
15.2973 | -34.7616 | -24.4615
15.3034 | -34.7621 | -24.4693
15.3095 | -34.7626 | -24.4772
15.3156 | -34.7631 | -24.485
15.3217 | -34.7636 | -24.4932
15.3278 | -34.7641 | -24.5015
15.3339 | -34.7646 | -24.5095
15.34 | -34.7651 | -24.5175
15.3461 | -34.7656 | -24.5255
15.3522 | -34.766 | -24.5333
15.3583 | -34.7665 | -24.5414
15.3644 | -34.767 | -24.5494
15.3705 | -34.7675 | -24.5572
15.3766 | -34.7679 | -24.5656
15.3827 | -34.7684 | -24.5736
15.3888 | -34.7689 | -24.5817
15.3949 | -34.7694 | -24.5898
15.401 | -34.7699 | -24.5979
15.4072 | -34.7703 | -24.6058
15.4133 | -34.7708 | -24.6139
15.4194 | -34.7712 | -24.6219
15.4255 | -34.7717 | -24.6299
15.4316 | -34.7722 | -24.6379
15.4377 | -34.7727 | -24.646
15.4438 | -34.7732 | -24.654
15.4499 | -34.7737 | -24.662
15.456 | -34.7742 | -24.6699
15.4621 | -34.7747 | -24.6778
15.4682 | -34.7752 | -24.6859
15.4743 | -34.7757 | -24.6936
15.4804 | -34.7761 | -24.7017
15.4865 | -34.7766 | -24.7096
15.4926 | -34.7771 | -24.7177
15.4987 | -34.7776 | -24.7258
15.5048 | -34.7781 | -24.7338
15.5109 | -34.7786 | -24.7418
15.517 | -34.7791 | -24.7499
15.5231 | -34.7796 | -24.758
15.5292 | -34.78 | -24.7662
15.5353 | -34.7805 | -24.7745
15.5414 | -34.781 | -24.7823
15.5476 | -34.7816 | -24.7901
15.5537 | -34.782 | -24.7982
15.5598 | -34.7826 | -24.8064
15.5659 | -34.783 | -24.8144
15.572 | -34.7835 | -24.8223
15.5781 | -34.784 | -24.8304
15.5842 | -34.7845 | -24.8384
15.5903 | -34.785 | -24.8461
15.5964 | -34.7855 | -24.8541
15.6025 | -34.786 | -24.862
15.6086 | -34.7864 | -24.8701
15.6147 | -34.7869 | -24.8781
15.6208 | -34.7874 | -24.886
15.6269 | -34.7878 | -24.894
15.633 | -34.7883 | -24.9021
15.6391 | -34.7889 | -24.9102
15.6452 | -34.7894 | -24.9183
15.6513 | -34.7898 | -24.926
15.6574 | -34.7903 | -24.9341
15.6635 | -34.7908 | -24.942
15.6696 | -34.7913 | -24.9501
15.6757 | -34.7918 | -24.958
15.6818 | -34.7923 | -24.9661
15.688 | -34.7928 | -24.974
15.6941 | -34.7933 | -24.982
15.7002 | -34.7938 | -24.9899
15.7063 | -34.7943 | -24.9979
15.7124 | -34.7948 | -25.0061
15.7185 | -34.7953 | -25.0141
15.7246 | -34.7958 | -25.022
15.7307 | -34.7963 | -25.03
15.7368 | -34.7968 | -25.0379
15.7429 | -34.7973 | -25.0461
15.749 | -34.7977 | -25.0541
15.7551 | -34.7982 | -25.0621
15.7612 | -34.7987 | -25.0701
15.7673 | -34.7992 | -25.078
15.7734 | -34.7998 | -25.0862
15.7795 | -34.8003 | -25.0944
15.7856 | -34.8008 | -25.1024
15.7917 | -34.8013 | -25.1101
15.7978 | -34.8017 | -25.1183
15.8039 | -34.8022 | -25.1263
15.81 | -34.8027 | -25.1342
15.8161 | -34.8032 | -25.1419
15.8222 | -34.8037 | -25.1498
15.8283 | -34.8042 | -25.1575
15.8345 | -34.8047 | -25.1657
15.8406 | -34.8052 | -25.1735
15.8467 | -34.8057 | -25.1817
15.8528 | -34.8062 | -25.1896
15.8589 | -34.8067 | -25.1975
15.865 | -34.8072 | -25.2052
15.8711 | -34.8077 | -25.2133
15.8772 | -34.8082 | -25.2211
15.8833 | -34.8087 | -25.2291
15.8894 | -34.8092 | -25.2371
15.8955 | -34.8097 | -25.2451
15.9016 | -34.8102 | -25.253
15.9077 | -34.8107 | -25.2608
15.9138 | -34.8111 | -25.269
15.9199 | -34.8116 | -25.2768
15.926 | -34.8121 | -25.2847
15.9321 | -34.8126 | -25.2924
15.9382 | -34.813 | -25.3003
15.9443 | -34.8135 | -25.3084
15.9504 | -34.814 | -25.3164
15.9565 | -34.8145 | -25.3242
15.9626 | -34.815 | -25.3321
15.9687 | -34.8155 | -25.3401
15.9749 | -34.816 | -25.348
15.981 | -34.8165 | -25.3558
15.9871 | -34.817 | -25.3638
15.9932 | -34.8174 | -25.3718
15.9993 | -34.818 | -25.3795
16.0054 | -34.8185 | -25.3873
16.0115 | -34.819 | -25.3954
16.0176 | -34.8195 | -25.4034
16.0237 | -34.82 | -25.4112
16.0298 | -34.8205 | -25.4189
16.0359 | -34.8209 | -25.4267
16.042 | -34.8214 | -25.4346
16.0481 | -34.8219 | -25.4426
16.0542 | -34.8224 | -25.4506
16.0603 | -34.8229 | -25.4584
16.0664 | -34.8234 | -25.4663
16.0725 | -34.8239 | -25.4742
16.0786 | -34.8244 | -25.4819
16.0847 | -34.8249 | -25.4898
16.0908 | -34.8254 | -25.4978
16.0969 | -34.8258 | -25.5058
16.103 | -34.8263 | -25.5136
16.1091 | -34.8268 | -25.5215
16.1152 | -34.8273 | -25.5292
16.1214 | -34.8278 | -25.5371
16.1275 | -34.8283 | -25.545
16.1336 | -34.8288 | -25.5529
16.1397 | -34.8293 | -25.5606
16.1458 | -34.8297 | -25.5683
16.1519 | -34.8302 | -25.5764
16.158 | -34.8307 | -25.5844
16.1641 | -34.8311 | -25.5923
16.1702 | -34.8316 | -25.6002
16.1763 | -34.8321 | -25.6079
16.1824 | -34.8326 | -25.6156
16.1885 | -34.833 | -25.6233
16.1946 | -34.8335 | -25.6309
16.2007 | -34.834 | -25.6389
16.2068 | -34.8345 | -25.6467
16.2129 | -34.835 | -25.6547
16.219 | -34.8356 | -25.6622
16.2251 | -34.8361 | -25.6697
16.2312 | -34.8366 | -25.6775
16.2373 | -34.837 | -25.6853
16.2434 | -34.8375 | -25.693
16.2495 | -34.838 | -25.7008
16.2556 | -34.8385 | -25.7084
16.2618 | -34.839 | -25.7161
16.2679 | -34.8395 | -25.724
16.274 | -34.84 | -25.7318
16.2801 | -34.8405 | -25.7397
16.2862 | -34.841 | -25.7471
16.2923 | -34.8415 | -25.7551
16.2984 | -34.842 | -25.7628
16.3045 | -34.8424 | -25.7707
16.3106 | -34.8429 | -25.7785
16.3167 | -34.8434 | -25.7864
16.3228 | -34.8439 | -25.7939
16.3289 | -34.8443 | -25.8017
16.335 | -34.8448 | -25.8097
16.3411 | -34.8453 | -25.8173
16.3472 | -34.8458 | -25.825
16.3533 | -34.8463 | -25.833
16.3594 | -34.8468 | -25.8406
16.3655 | -34.8473 | -25.8483
16.3716 | -34.8478 | -25.8562
16.3777 | -34.8483 | -25.8638
16.3838 | -34.8488 | -25.8714
16.3899 | -34.8493 | -25.8791
16.396 | -34.8498 | -25.8868
16.4021 | -34.8503 | -25.8945
16.4083 | -34.8508 | -25.9023
16.4144 | -34.8513 | -25.9102
16.4205 | -34.8518 | -25.9179
16.4266 | -34.8523 | -25.9255
16.4327 | -34.8527 | -25.9332
16.4388 | -34.8532 | -25.941
16.4449 | -34.8537 | -25.9486
16.451 | -34.8541 | -25.9562
16.4571 | -34.8546 | -25.964
16.4632 | -34.8551 | -25.9717
16.4693 | -34.8556 | -25.9795
16.4754 | -34.8561 | -25.9874
16.4815 | -34.8566 | -25.9951
16.4876 | -34.8571 | -26.0028
16.4937 | -34.8575 | -26.0105
16.4998 | -34.858 | -26.0179
16.5059 | -34.8585 | -26.0255
16.512 | -34.8589 | -26.0331
16.5181 | -34.8595 | -26.0411
16.5242 | -34.8599 | -26.0491
16.5303 | -34.8604 | -26.0566
16.5364 | -34.8609 | -26.0644
16.5425 | -34.8614 | -26.0721
16.5487 | -34.8618 | -26.0797
16.5548 | -34.8623 | -26.0876
16.5609 | -34.8629 | -26.0954
16.567 | -34.8633 | -26.1024
16.5731 | -34.8638 | -26.1104
16.5792 | -34.8643 | -26.1181
16.5853 | -34.8648 | -26.1258
16.5914 | -34.8652 | -26.1336
16.5975 | -34.8657 | -26.1416
16.6036 | -34.8662 | -26.149
16.6097 | -34.8667 | -26.1568
16.6158 | -34.8672 | -26.1644
16.6219 | -34.8677 | -26.1722
16.628 | -34.8681 | -26.1799
16.6341 | -34.8686 | -26.1874
16.6402 | -34.8691 | -26.1948
16.6463 | -34.8695 | -26.2025
16.6524 | -34.87 | -26.21
16.6585 | -34.8705 | -26.2177
16.6646 | -34.871 | -26.2253
16.6707 | -34.8715 | -26.2325
16.6768 | -34.872 | -26.2403
16.6829 | -34.8725 | -26.2481
16.689 | -34.8729 | -26.2558
16.6952 | -34.8734 | -26.2635
16.7013 | -34.8739 | -26.2712
16.7074 | -34.8744 | -26.2786
16.7135 | -34.8749 | -26.2862
16.7196 | -34.8754 | -26.294
16.7257 | -34.8759 | -26.3017
16.7318 | -34.8764 | -26.3096
16.7379 | -34.8769 | -26.3171
16.744 | -34.8773 | -26.3247
16.7501 | -34.8779 | -26.3325
16.7562 | -34.8783 | -26.3401
16.7623 | -34.8788 | -26.3478
16.7684 | -34.8793 | -26.3556
16.7745 | -34.8798 | -26.3634
16.7806 | -34.8803 | -26.3708
16.7867 | -34.8808 | -26.3786
16.7928 | -34.8813 | -26.3863
16.7989 | -34.8818 | -26.394
16.805 | -34.8823 | -26.4018
16.8111 | -34.8827 | -26.4092
16.8172 | -34.8832 | -26.417
16.8233 | -34.8837 | -26.425
16.8294 | -34.8842 | -26.4325
16.8356 | -34.8847 | -26.44
16.8417 | -34.8852 | -26.4474
16.8478 | -34.8857 | -26.4549
16.8539 | -34.8862 | -26.4626
16.86 | -34.8866 | -26.4705
16.8661 | -34.8871 | -26.4781
16.8722 | -34.8876 | -26.4859
16.8783 | -34.8881 | -26.4937
16.8844 | -34.8886 | -26.5011
16.8905 | -34.8891 | -26.5086
16.8966 | -34.8896 | -26.5162
16.9027 | -34.8901 | -26.5238
16.9088 | -34.8905 | -26.5314
16.9149 | -34.891 | -26.5391
16.921 | -34.8915 | -26.5464
16.9271 | -34.8919 | -26.5543
16.9332 | -34.8924 | -26.5617
16.9393 | -34.8929 | -26.5691
16.9454 | -34.8933 | -26.5766
16.9515 | -34.8938 | -26.5844
16.9576 | -34.8942 | -26.5916
16.9637 | -34.8947 | -26.5991
16.9698 | -34.8952 | -26.6066
16.9759 | -34.8957 | -26.6141
16.9821 | -34.8961 | -26.6217
16.9882 | -34.8967 | -26.6286
16.9943 | -34.8971 | -26.6361
17.0004 | -34.8976 | -26.6436
17.0065 | -34.8981 | -26.6513
17.0126 | -34.8986 | -26.6587
17.0187 | -34.8991 | -26.6663
17.0248 | -34.8996 | -26.6736
17.0309 | -34.9002 | -26.6811
17.037 | -34.9006 | -26.6886
17.0431 | -34.9011 | -26.6964
17.0492 | -34.9016 | -26.7041
17.0553 | -34.9021 | -26.7117
17.0614 | -34.9026 | -26.7192
17.0675 | -34.9031 | -26.7265
17.0736 | -34.9036 | -26.7341
17.0797 | -34.9041 | -26.7417
17.0858 | -34.9046 | -26.7494
17.0919 | -34.905 | -26.757
17.098 | -34.9055 | -26.7642
17.1041 | -34.9059 | -26.7714
17.1102 | -34.9064 | -26.7789
17.1163 | -34.907 | -26.7865
17.1225 | -34.9075 | -26.7938
17.1286 | -34.9079 | -26.801
17.1347 | -34.9084 | -26.8083
17.1408 | -34.9088 | -26.8159
17.1469 | -34.9093 | -26.8231
17.153 | -34.9098 | -26.8306
17.1591 | -34.9103 | -26.838
17.1652 | -34.9108 | -26.8451
17.1713 | -34.9112 | -26.8526
17.1774 | -34.9118 | -26.8601
17.1835 | -34.9123 | -26.8675
17.1896 | -34.9128 | -26.875
17.1957 | -34.9132 | -26.8825
17.2018 | -34.9137 | -26.8897
17.2079 | -34.9142 | -26.8975
17.214 | -34.9146 | -26.9048
17.2201 | -34.9151 | -26.9123
17.2262 | -34.9156 | -26.92
17.2323 | -34.9161 | -26.9276
17.2384 | -34.9166 | -26.935
17.2445 | -34.917 | -26.9426
17.2506 | -34.9175 | -26.95
17.2567 | -34.918 | -26.9575
17.2628 | -34.9184 | -26.9646
17.269 | -34.9189 | -26.9721
17.2751 | -34.9194 | -26.9793
17.2812 | -34.9198 | -26.9868
17.2873 | -34.9203 | -26.9943
17.2934 | -34.9208 | -27.0017
17.2995 | -34.9213 | -27.0094
17.3056 | -34.9217 | -27.0167
17.3117 | -34.9222 | -27.0242
17.3178 | -34.9227 | -27.0316
17.3239 | -34.9232 | -27.0391
17.33 | -34.9236 | -27.0464
17.3361 | -34.9241 | -27.0539
17.3422 | -34.9246 | -27.0611
17.3483 | -34.9251 | -27.0684
17.3544 | -34.9255 | -27.0761
17.3605 | -34.926 | -27.0836
17.3666 | -34.9265 | -27.091
17.3727 | -34.9269 | -27.0983
17.3788 | -34.9274 | -27.1054
17.3849 | -34.9279 | -27.1127
17.391 | -34.9284 | -27.1202
17.3971 | -34.9288 | -27.1276
17.4032 | -34.9293 | -27.1349
17.4094 | -34.9298 | -27.1421
17.4155 | -34.9303 | -27.1493
17.4216 | -34.9308 | -27.1567
17.4277 | -34.9312 | -27.1642
17.4338 | -34.9317 | -27.1713
17.4399 | -34.9322 | -27.1788
17.446 | -34.9326 | -27.1857
17.4521 | -34.9331 | -27.1933
17.4582 | -34.9336 | -27.2005
17.4643 | -34.934 | -27.2077
17.4704 | -34.9345 | -27.2149
17.4765 | -34.935 | -27.2223
17.4826 | -34.9355 | -27.2295
17.4887 | -34.936 | -27.237
17.4948 | -34.9364 | -27.2442
17.5009 | -34.9369 | -27.2515
17.507 | -34.9374 | -27.2589
17.5131 | -34.9379 | -27.2663
17.5192 | -34.9383 | -27.2734
17.5253 | -34.9388 | -27.2808
17.5314 | -34.9393 | -27.288
17.5375 | -34.9397 | -27.2955
17.5436 | -34.9402 | -27.3029
17.5497 | -34.9407 | -27.3102
17.5559 | -34.9412 | -27.3174
17.562 | -34.9417 | -27.3247
17.5681 | -34.9422 | -27.332
17.5742 | -34.9426 | -27.3396
17.5803 | -34.9431 | -27.3471
17.5864 | -34.9436 | -27.3541
17.5925 | -34.9441 | -27.3616
17.5986 | -34.9445 | -27.3687
17.6047 | -34.945 | -27.3762
17.6108 | -34.9455 | -27.3833
17.6169 | -34.946 | -27.3905
17.623 | -34.9465 | -27.3978
17.6291 | -34.9469 | -27.4051
17.6352 | -34.9475 | -27.4124
17.6413 | -34.9479 | -27.4199
17.6474 | -34.9484 | -27.4272
17.6535 | -34.9489 | -27.4345
17.6596 | -34.9494 | -27.4419
17.6657 | -34.9499 | -27.4493
17.6718 | -34.9504 | -27.4566
17.6779 | -34.9508 | -27.4639
17.684 | -34.9513 | -27.4713
17.6901 | -34.9518 | -27.4787
17.6963 | -34.9523 | -27.4857
17.7024 | -34.9527 | -27.4928
17.7085 | -34.9532 | -27.4999
17.7146 | -34.9537 | -27.5073
17.7207 | -34.9542 | -27.5145
17.7268 | -34.9547 | -27.5216
17.7329 | -34.9552 | -27.5289
17.739 | -34.9557 | -27.5364
17.7451 | -34.9562 | -27.5435
17.7512 | -34.9566 | -27.5509
17.7573 | -34.9571 | -27.5584
17.7634 | -34.9576 | -27.5655
17.7695 | -34.9581 | -27.5727
17.7756 | -34.9586 | -27.5799
17.7817 | -34.959 | -27.5871
17.7878 | -34.9595 | -27.5943
17.7939 | -34.96 | -27.6016
17.8 | -34.9605 | -27.6092
17.8061 | -34.9609 | -27.6165
17.8122 | -34.9614 | -27.6238
17.8183 | -34.9619 | -27.631
17.8244 | -34.9624 | -27.6384
17.8305 | -34.9628 | -27.6456
17.8366 | -34.9633 | -27.6527
17.8428 | -34.9638 | -27.6597
17.8489 | -34.9643 | -27.667
17.855 | -34.9648 | -27.6743
17.8611 | -34.9653 | -27.6817
17.8672 | -34.9657 | -27.6892
17.8733 | -34.9662 | -27.6966
17.8794 | -34.9667 | -27.704
17.8855 | -34.9672 | -27.7114
17.8916 | -34.9676 | -27.7187
17.8977 | -34.9681 | -27.726
17.9038 | -34.9686 | -27.7327
17.9099 | -34.9691 | -27.7398
17.916 | -34.9695 | -27.7472
17.9221 | -34.97 | -27.7545
17.9282 | -34.9705 | -27.7616
17.9343 | -34.971 | -27.7689
17.9404 | -34.9715 | -27.7758
17.9465 | -34.9719 | -27.7827
17.9526 | -34.9725 | -27.7899
17.9587 | -34.9729 | -27.797
17.9648 | -34.9734 | -27.8042
17.9709 | -34.9739 | -27.8114
17.977 | -34.9744 | -27.8185
17.9832 | -34.9749 | -27.8257
17.9893 | -34.9754 | -27.8331
17.9954 | -34.9759 | -27.8405
18.0015 | -34.9764 | -27.8479
18.0076 | -34.9769 | -27.8548
18.0137 | -34.9773 | -27.8622
18.0198 | -34.9778 | -27.8693
18.0259 | -34.9782 | -27.8766
18.032 | -34.9786 | -27.8839
18.0381 | -34.9791 | -27.8913
18.0442 | -34.9797 | -27.8982
18.0503 | -34.9801 | -27.9054
18.0564 | -34.9806 | -27.9127
18.0625 | -34.9811 | -27.9201
18.0686 | -34.9815 | -27.9272
18.0747 | -34.982 | -27.9344
18.0808 | -34.9825 | -27.9416
18.0869 | -34.983 | -27.9487
18.093 | -34.9835 | -27.9557
18.0991 | -34.984 | -27.9625
18.1052 | -34.9845 | -27.9696
18.1113 | -34.9849 | -27.9766
18.1174 | -34.9854 | -27.9835
18.1236 | -34.9859 | -27.991
18.1297 | -34.9864 | -27.998
18.1358 | -34.9869 | -28.0049
18.1419 | -34.9874 | -28.0124
18.148 | -34.9879 | -28.0194
18.1541 | -34.9884 | -28.0267
18.1602 | -34.9889 | -28.0338
18.1663 | -34.9894 | -28.0407
18.1724 | -34.9898 | -28.048
18.1785 | -34.9903 | -28.0549
18.1846 | -34.9908 | -28.0619
18.1907 | -34.9913 | -28.0689
18.1968 | -34.9917 | -28.0762
18.2029 | -34.9922 | -28.0833
18.209 | -34.9926 | -28.0906
18.2151 | -34.9931 | -28.0979
18.2212 | -34.9936 | -28.1048
18.2273 | -34.9941 | -28.112
18.2334 | -34.9946 | -28.119
18.2395 | -34.9951 | -28.1262
18.2456 | -34.9956 | -28.1335
18.2517 | -34.996 | -28.1407
18.2578 | -34.9965 | -28.1476
18.2639 | -34.9971 | -28.1547
18.2701 | -34.9975 | -28.162
18.2762 | -34.998 | -28.1693
18.2823 | -34.9984 | -28.1762
18.2884 | -34.9989 | -28.1836
18.2945 | -34.9994 | -28.1905
18.3006 | -34.9998 | -28.1976
18.3067 | -35.0003 | -28.2047
18.3128 | -35.0008 | -28.2119
18.3189 | -35.0013 | -28.2193
18.325 | -35.0017 | -28.2259
18.3311 | -35.0022 | -28.2331
18.3372 | -35.0027 | -28.2403
18.3433 | -35.0032 | -28.2477
18.3494 | -35.0036 | -28.2548
18.3555 | -35.0041 | -28.262
18.3616 | -35.0046 | -28.2691
18.3677 | -35.0051 | -28.2761
18.3738 | -35.0056 | -28.2833
18.3799 | -35.0061 | -28.2904
18.386 | -35.0066 | -28.2977
18.3921 | -35.0071 | -28.3049
18.3982 | -35.0075 | -28.312
18.4043 | -35.008 | -28.3193
18.4105 | -35.0085 | -28.3266
18.4166 | -35.0089 | -28.3335
18.4227 | -35.0094 | -28.3407
18.4288 | -35.0099 | -28.3479
18.4349 | -35.0104 | -28.355
18.441 | -35.0108 | -28.3621
18.4471 | -35.0113 | -28.3694
18.4532 | -35.0118 | -28.3768
18.4593 | -35.0123 | -28.384
18.4654 | -35.0127 | -28.3907
18.4715 | -35.0132 | -28.3976
18.4776 | -35.0138 | -28.4049
18.4837 | -35.0142 | -28.412
18.4898 | -35.0147 | -28.4191
18.4959 | -35.0152 | -28.4265
18.502 | -35.0157 | -28.4335
18.5081 | -35.0162 | -28.4407
18.5142 | -35.0166 | -28.4478
18.5203 | -35.0171 | -28.4546
18.5264 | -35.0176 | -28.4617
18.5325 | -35.0181 | -28.4688
18.5386 | -35.0186 | -28.4757
18.5447 | -35.0191 | -28.4826
18.5508 | -35.0196 | -28.4898
18.557 | -35.0201 | -28.4969
18.5631 | -35.0205 | -28.5041
18.5692 | -35.021 | -28.5112
18.5753 | -35.0215 | -28.5182
18.5814 | -35.0219 | -28.5253
18.5875 | -35.0224 | -28.5324
18.5936 | -35.0229 | -28.5396
18.5997 | -35.0233 | -28.5468
18.6058 | -35.0238 | -28.5539
18.6119 | -35.0243 | -28.5611
18.618 | -35.0248 | -28.5683
18.6241 | -35.0252 | -28.5755
18.6302 | -35.0257 | -28.5826
18.6363 | -35.0262 | -28.5897
18.6424 | -35.0267 | -28.5967
18.6485 | -35.0272 | -28.6038
18.6546 | -35.0277 | -28.6109
18.6607 | -35.0281 | -28.6182
18.6668 | -35.0286 | -28.6252
18.6729 | -35.0291 | -28.6325
18.679 | -35.0296 | -28.6392
18.6851 | -35.0301 | -28.6463
18.6912 | -35.0306 | -28.6533
18.6974 | -35.0311 | -28.6604
18.7035 | -35.0315 | -28.6675
18.7096 | -35.032 | -28.675
18.7157 | -35.0325 | -28.6819
18.7218 | -35.033 | -28.689
18.7279 | -35.0335 | -28.6962
18.734 | -35.0339 | -28.7033
18.7401 | -35.0344 | -28.7105
18.7462 | -35.0348 | -28.7173
18.7523 | -35.0353 | -28.7245
18.7584 | -35.0358 | -28.7319
18.7645 | -35.0363 | -28.7387
18.7706 | -35.0367 | -28.7458
18.7767 | -35.0372 | -28.7529
18.7828 | -35.0376 | -28.7598
18.7889 | -35.0381 | -28.7666
18.795 | -35.0386 | -28.774
18.8011 | -35.0391 | -28.7811
18.8072 | -35.0396 | -28.7883
18.8133 | -35.04 | -28.7955
18.8194 | -35.0405 | -28.8025
18.8255 | -35.041 | -28.8096
18.8316 | -35.0414 | -28.8168
18.8377 | -35.0419 | -28.8235
18.8439 | -35.0424 | -28.8305
18.85 | -35.0429 | -28.8376
18.8561 | -35.0434 | -28.8447
18.8622 | -35.0439 | -28.852
18.8683 | -35.0444 | -28.8593
18.8744 | -35.0449 | -28.8663
18.8805 | -35.0454 | -28.8732
18.8866 | -35.0459 | -28.8805
18.8927 | -35.0464 | -28.8872
18.8988 | -35.0469 | -28.8941
18.9049 | -35.0473 | -28.9009
18.911 | -35.0478 | -28.9077
18.9171 | -35.0483 | -28.9149
18.9232 | -35.0488 | -28.9218
18.9293 | -35.0492 | -28.9289
18.9354 | -35.0497 | -28.9361
18.9415 | -35.0502 | -28.9429
18.9476 | -35.0506 | -28.9498
18.9537 | -35.0511 | -28.9569
18.9598 | -35.0515 | -28.9637
18.9659 | -35.052 | -28.9707
18.972 | -35.0525 | -28.9777
18.9781 | -35.053 | -28.9848
18.9843 | -35.0535 | -28.992
18.9904 | -35.0539 | -28.9993
18.9965 | -35.0544 | -29.0059
19.0026 | -35.0549 | -29.013
19.0087 | -35.0554 | -29.0204
19.0148 | -35.0559 | -29.0275
19.0209 | -35.0564 | -29.0345
19.027 | -35.0569 | -29.0418
19.0331 | -35.0574 | -29.0487
19.0392 | -35.0579 | -29.0558
19.0453 | -35.0584 | -29.0629
19.0514 | -35.0588 | -29.0696
19.0575 | -35.0593 | -29.0769
19.0636 | -35.0598 | -29.0838
19.0697 | -35.0603 | -29.091
19.0758 | -35.0608 | -29.0981
19.0819 | -35.0613 | -29.105
19.088 | -35.0618 | -29.112
19.0941 | -35.0624 | -29.1191
19.1002 | -35.0628 | -29.1259
19.1063 | -35.0633 | -29.1326
19.1124 | -35.0638 | -29.1398
19.1185 | -35.0642 | -29.1467
19.1246 | -35.0647 | -29.1537
19.1308 | -35.0652 | -29.1605
19.1369 | -35.0657 | -29.1677
19.143 | -35.0662 | -29.1748
19.1491 | -35.0667 | -29.1819
19.1552 | -35.0671 | -29.1893
19.1613 | -35.0676 | -29.1962
19.1674 | -35.068 | -29.2036
19.1735 | -35.0685 | -29.2107
19.1796 | -35.069 | -29.2175
19.1857 | -35.0695 | -29.2245
19.1918 | -35.07 | -29.2318
19.1979 | -35.0704 | -29.2387
19.204 | -35.0709 | -29.2459
19.2101 | -35.0713 | -29.2529
19.2162 | -35.0718 | -29.2596
19.2223 | -35.0722 | -29.2668
19.2284 | -35.0727 | -29.2741
19.2345 | -35.0732 | -29.2811
19.2406 | -35.0737 | -29.2881
19.2467 | -35.0742 | -29.2953
19.2528 | -35.0746 | -29.3024
19.2589 | -35.0751 | -29.3091
19.265 | -35.0756 | -29.3163
19.2712 | -35.076 | -29.3235
19.2773 | -35.0765 | -29.3307
19.2834 | -35.077 | -29.3378
19.2895 | -35.0775 | -29.3451
19.2956 | -35.078 | -29.3522
19.3017 | -35.0784 | -29.359
19.3078 | -35.0789 | -29.3661
19.3139 | -35.0793 | -29.3729
19.32 | -35.0797 | -29.3798
19.3261 | -35.0802 | -29.3867
19.3322 | -35.0807 | -29.3938
19.3383 | -35.0812 | -29.4008
19.3444 | -35.0816 | -29.4078
19.3505 | -35.0821 | -29.4144
19.3566 | -35.0826 | -29.4215
19.3627 | -35.0831 | -29.4288
19.3688 | -35.0835 | -29.4353
19.3749 | -35.084 | -29.442
19.381 | -35.0845 | -29.4486
19.3871 | -35.085 | -29.4557
19.3932 | -35.0854 | -29.4627
19.3993 | -35.0859 | -29.4697
19.4054 | -35.0864 | -29.4768
19.4115 | -35.0869 | -29.4839
19.4177 | -35.0873 | -29.4909
19.4238 | -35.0878 | -29.4976
19.4299 | -35.0883 | -29.5049
19.436 | -35.0887 | -29.5121
19.4421 | -35.0892 | -29.519
19.4482 | -35.0897 | -29.5257
19.4543 | -35.0901 | -29.5325
19.4604 | -35.0906 | -29.5394
19.4665 | -35.0911 | -29.5464
19.4726 | -35.0916 | -29.5529
19.4787 | -35.0921 | -29.5601
19.4848 | -35.0926 | -29.5672
19.4909 | -35.0931 | -29.5742
19.497 | -35.0935 | -29.5813
19.5031 | -35.094 | -29.5881
19.5092 | -35.0944 | -29.5952
19.5153 | -35.0949 | -29.6019
19.5214 | -35.0953 | -29.6085
19.5275 | -35.0957 | -29.6156
19.5336 | -35.0961 | -29.6227
19.5397 | -35.0965 | -29.6297
19.5458 | -35.097 | -29.6369
19.5519 | -35.0975 | -29.6437
19.5581 | -35.098 | -29.6505
19.5642 | -35.0985 | -29.6575
19.5703 | -35.0989 | -29.665
19.5764 | -35.0994 | -29.6723
19.5825 | -35.0998 | -29.6789
19.5886 | -35.1003 | -29.6857
19.5947 | -35.1007 | -29.6923
19.6008 | -35.1012 | -29.6993
19.6069 | -35.1017 | -29.706
19.613 | -35.1022 | -29.7131
19.6191 | -35.1026 | -29.72
19.6252 | -35.1031 | -29.727
19.6313 | -35.1036 | -29.7338
19.6374 | -35.1041 | -29.7407
19.6435 | -35.1046 | -29.748
19.6496 | -35.105 | -29.7548
19.6557 | -35.1056 | -29.7619
19.6618 | -35.1061 | -29.7689
19.6679 | -35.1066 | -29.7759
19.674 | -35.107 | -29.7828
19.6801 | -35.1075 | -29.7896
19.6862 | -35.1079 | -29.7967
19.6923 | -35.1083 | -29.8039
19.6984 | -35.1088 | -29.8107
19.7046 | -35.1093 | -29.8183
19.7107 | -35.1097 | -29.8255
19.7168 | -35.1102 | -29.8325
19.7229 | -35.1107 | -29.8396
19.729 | -35.1111 | -29.8464
19.7351 | -35.1117 | -29.8531
19.7412 | -35.1122 | -29.8603
19.7473 | -35.1126 | -29.8674
19.7534 | -35.1131 | -29.8743
19.7595 | -35.1136 | -29.8808
19.7656 | -35.114 | -29.8877
19.7717 | -35.1145 | -29.8945
19.7778 | -35.115 | -29.9012
19.7839 | -35.1154 | -29.9081
19.79 | -35.1159 | -29.915
19.7961 | -35.1163 | -29.922
19.8022 | -35.1168 | -29.9286
19.8083 | -35.1173 | -29.9357
19.8144 | -35.1178 | -29.9426
19.8205 | -35.1183 | -29.9494
19.8266 | -35.1188 | -29.956
19.8327 | -35.1193 | -29.9627
19.8388 | -35.1197 | -29.9687
19.845 | -35.1201 | -29.9757
19.8511 | -35.1206 | -29.9828
19.8572 | -35.121 | -29.9898
19.8633 | -35.1214 | -29.9967
19.8694 | -35.1219 | -30.0035
19.8755 | -35.1225 | -30.0109
19.8816 | -35.1229 | -30.018
19.8877 | -35.1233 | -30.0248
19.8938 | -35.1238 | -30.0317
19.8999 | -35.1243 | -30.039
19.906 | -35.1247 | -30.0462
19.9121 | -35.1254 | -30.0531
19.9182 | -35.1258 | -30.0601
19.9243 | -35.1263 | -30.0673
19.9304 | -35.1267 | -30.074
19.9365 | -35.1273 | -30.0811
19.9426 | -35.1277 | -30.0877
19.9487 | -35.1282 | -30.0945
19.9548 | -35.1286 | -30.1016
19.9609 | -35.1291 | -30.1084
19.967 | -35.1296 | -30.115
19.9731 | -35.1301 | -30.1222
19.9792 | -35.1305 | -30.1293
19.9853 | -35.131 | -30.1359
19.9915 | -35.1315 | -30.1419
19.9976 | -35.132 | -30.1487
20.0037 | -35.1324 | -30.1557
20.0098 | -35.1329 | -30.1628
20.0159 | -35.1335 | -30.17
20.022 | -35.134 | -30.1768
20.0281 | -35.1344 | -30.1844
20.0342 | -35.1349 | -30.1914
20.0403 | -35.1354 | -30.1986
20.0464 | -35.136 | -30.2062
20.0525 | -35.1365 | -30.2125
20.0586 | -35.1369 | -30.2192
20.0647 | -35.1374 | -30.2261
20.0708 | -35.138 | -30.2326
20.0769 | -35.1384 | -30.239
20.083 | -35.1389 | -30.2451
20.0891 | -35.1393 | -30.2517
20.0952 | -35.1397 | -30.2583
20.1013 | -35.1402 | -30.2651
20.1074 | -35.1406 | -30.2723
20.1135 | -35.1411 | -30.2788
20.1196 | -35.1416 | -30.2847
20.1257 | -35.1421 | -30.2914
20.1319 | -35.1426 | -30.2977
20.138 | -35.143 | -30.3046
20.1441 | -35.1435 | -30.312
20.1502 | -35.1439 | -30.3187
20.1563 | -35.1444 | -30.3249
20.1624 | -35.1448 | -30.3317
20.1685 | -35.1452 | -30.3388
20.1746 | -35.1457 | -30.3462
20.1807 | -35.1462 | -30.3531
20.1868 | -35.1466 | -30.3595
20.1929 | -35.1471 | -30.3664
20.199 | -35.1475 | -30.3735
20.2051 | -35.1479 | -30.3801
20.2112 | -35.1484 | -30.3869
20.2173 | -35.1489 | -30.3935
20.2234 | -35.1493 | -30.4004
20.2295 | -35.1497 | -30.4076
20.2356 | -35.1502 | -30.4145
20.2417 | -35.1506 | -30.4219
20.2478 | -35.1509 | -30.4289
20.2539 | -35.1515 | -30.4362
20.26 | -35.1519 | -30.4429
20.2661 | -35.1523 | -30.4495
20.2723 | -35.1528 | -30.4559
20.2784 | -35.1533 | -30.462
20.2845 | -35.1538 | -30.4685
20.2906 | -35.1543 | -30.4749
20.2967 | -35.1548 | -30.4818
20.3028 | -35.1553 | -30.4885
20.3089 | -35.1558 | -30.4953
20.315 | -35.1563 | -30.5022
20.3211 | -35.1567 | -30.5091
20.3272 | -35.1572 | -30.5153
20.3333 | -35.1576 | -30.5212
20.3394 | -35.1581 | -30.5278
20.3455 | -35.1587 | -30.5351
20.3516 | -35.1592 | -30.5414
20.3577 | -35.1597 | -30.5477
20.3638 | -35.1601 | -30.5552
20.3699 | -35.1605 | -30.5613
20.376 | -35.161 | -30.568
20.3821 | -35.1615 | -30.575
20.3882 | -35.1619 | -30.5816
20.3943 | -35.1624 | -30.5885
20.4004 | -35.1627 | -30.5953
20.4065 | -35.1632 | -30.6024
20.4126 | -35.1637 | -30.6089
20.4188 | -35.1641 | -30.616
20.4249 | -35.1645 | -30.6231
20.431 | -35.165 | -30.6298
20.4371 | -35.1653 | -30.6354
20.4432 | -35.1658 | -30.6417
20.4493 | -35.1663 | -30.6487
20.4554 | -35.1668 | -30.6557
20.4615 | -35.1672 | -30.6627
20.4676 | -35.1677 | -30.6693
20.4737 | -35.1681 | -30.6758
20.4798 | -35.1687 | -30.6826
20.4859 | -35.1692 | -30.6902
20.492 | -35.1697 | -30.6965
20.4981 | -35.1701 | -30.7025
20.5042 | -35.1706 | -30.7096
20.5103 | -35.171 | -30.7161
20.5164 | -35.1715 | -30.723
20.5225 | -35.1719 | -30.7297
20.5286 | -35.1723 | -30.7368
20.5347 | -35.1727 | -30.7436
20.5408 | -35.1731 | -30.7502
20.5469 | -35.1735 | -30.7576
20.553 | -35.1739 | -30.7643
20.5592 | -35.1744 | -30.7709
20.5653 | -35.1749 | -30.7775
20.5714 | -35.1753 | -30.7839
20.5775 | -35.1758 | -30.7902
20.5836 | -35.1761 | -30.7969
20.5897 | -35.1766 | -30.8038
20.5958 | -35.1771 | -30.8098
20.6019 | -35.1776 | -30.8167
20.608 | -35.178 | -30.8237
20.6141 | -35.1784 | -30.8305
20.6202 | -35.1788 | -30.8373
20.6263 | -35.1793 | -30.844
20.6324 | -35.1798 | -30.8505
20.6385 | -35.1803 | -30.8569
20.6446 | -35.1807 | -30.8634
20.6507 | -35.1811 | -30.8704
20.6568 | -35.1815 | -30.877
20.6629 | -35.1819 | -30.8838
20.669 | -35.1823 | -30.8904
20.6751 | -35.1828 | -30.897
20.6812 | -35.1833 | -30.9032
20.6873 | -35.1838 | -30.9101
20.6934 | -35.1842 | -30.9167
20.6995 | -35.1847 | -30.9235
20.7057 | -35.1852 | -30.9303
20.7118 | -35.1858 | -30.9371
20.7179 | -35.1862 | -30.9436
20.724 | -35.1866 | -30.9506
20.7301 | -35.1871 | -30.9576
20.7362 | -35.1875 | -30.9641
20.7423 | -35.188 | -30.9706
20.7484 | -35.1884 | -30.9772
20.7545 | -35.1889 | -30.984
20.7606 | -35.1894 | -30.9909
20.7667 | -35.1899 | -30.9979
20.7728 | -35.1904 | -31.0045
20.7789 | -35.1908 | -31.0114
20.785 | -35.1912 | -31.0181
20.7911 | -35.1918 | -31.0243
20.7972 | -35.1923 | -31.0305
20.8033 | -35.1928 | -31.0372
20.8094 | -35.1933 | -31.0435
20.8155 | -35.1937 | -31.05
20.8216 | -35.1942 | -31.0566
20.8277 | -35.1947 | -31.0633
20.8338 | -35.1951 | -31.0697
20.8399 | -35.1956 | -31.0763
20.8461 | -35.196 | -31.0827
20.8522 | -35.1965 | -31.0893
20.8583 | -35.197 | -31.0965
20.8644 | -35.1975 | -31.1034
20.8705 | -35.1979 | -31.1102
20.8766 | -35.1984 | -31.1172
20.8827 | -35.1989 | -31.1238
20.8888 | -35.1993 | -31.1305
20.8949 | -35.1997 | -31.1371
20.901 | -35.2002 | -31.1439
20.9071 | -35.2006 | -31.1507
20.9132 | -35.2011 | -31.1571
20.9193 | -35.2014 | -31.1639
20.9254 | -35.2019 | -31.1706
20.9315 | -35.2023 | -31.1772
20.9376 | -35.2028 | -31.1838
20.9437 | -35.2032 | -31.1906
20.9498 | -35.2036 | -31.1971
20.9559 | -35.2041 | -31.2038
20.962 | -35.2045 | -31.2102
20.9681 | -35.205 | -31.2167
20.9742 | -35.2054 | -31.2231
20.9803 | -35.2059 | -31.2296
20.9864 | -35.2063 | -31.2365
20.9926 | -35.2068 | -31.2431
20.9987 | -35.2072 | -31.2491
21.0048 | -35.2077 | -31.2556
21.0109 | -35.2082 | -31.2623
21.017 | -35.2086 | -31.2687
21.0231 | -35.209 | -31.2752
21.0292 | -35.2095 | -31.2819
21.0353 | -35.2099 | -31.2878
21.0414 | -35.2104 | -31.2943
21.0475 | -35.2108 | -31.3013
21.0536 | -35.2113 | -31.3077
21.0597 | -35.2117 | -31.3147
21.0658 | -35.2122 | -31.3214
21.0719 | -35.2126 | -31.3286
21.078 | -35.2131 | -31.3353
21.0841 | -35.2135 | -31.3418
21.0902 | -35.214 | -31.3483
21.0963 | -35.2145 | -31.3544
21.1024 | -35.215 | -31.3613
21.1085 | -35.2155 | -31.3677
21.1146 | -35.2159 | -31.3745
21.1207 | -35.2164 | -31.3812
21.1268 | -35.2168 | -31.3882
21.133 | -35.2173 | -31.3949
21.1391 | -35.2177 | -31.4014
21.1452 | -35.2181 | -31.408
21.1513 | -35.2187 | -31.4143
21.1574 | -35.2191 | -31.4206
21.1635 | -35.2196 | -31.4272
21.1696 | -35.22 | -31.4336
21.1757 | -35.2205 | -31.4397
21.1818 | -35.221 | -31.4465
21.1879 | -35.2214 | -31.4528
21.194 | -35.2218 | -31.4597
21.2001 | -35.2224 | -31.4664
21.2062 | -35.2228 | -31.4728
21.2123 | -35.2233 | -31.4786
21.2184 | -35.2238 | -31.4852
21.2245 | -35.2242 | -31.4919
21.2306 | -35.2247 | -31.4985
21.2367 | -35.2251 | -31.5052
21.2428 | -35.2256 | -31.5112
21.2489 | -35.226 | -31.5174
21.255 | -35.2264 | -31.5242
21.2611 | -35.2269 | -31.5307
21.2672 | -35.2274 | -31.5373
21.2733 | -35.2278 | -31.5437
21.2795 | -35.2283 | -31.5502
21.2856 | -35.2287 | -31.5564
21.2917 | -35.2291 | -31.5631
21.2978 | -35.2295 | -31.5695
21.3039 | -35.2299 | -31.5759
21.31 | -35.2304 | -31.5824
21.3161 | -35.2308 | -31.5887
21.3222 | -35.2313 | -31.5951
21.3283 | -35.2317 | -31.6015
21.3344 | -35.2322 | -31.6077
21.3405 | -35.2327 | -31.6141
21.3466 | -35.2331 | -31.6205
21.3527 | -35.2335 | -31.6259
21.3588 | -35.234 | -31.6322
21.3649 | -35.2344 | -31.6386
21.371 | -35.2349 | -31.645
21.3771 | -35.2354 | -31.6515
21.3832 | -35.2358 | -31.6576
21.3893 | -35.2363 | -31.6632
21.3954 | -35.2368 | -31.6698
21.4015 | -35.2373 | -31.6763
21.4076 | -35.2377 | -31.6828
21.4137 | -35.2381 | -31.6894
21.4199 | -35.2386 | -31.6955
21.426 | -35.239 | -31.702
21.4321 | -35.2394 | -31.7089
21.4382 | -35.2399 | -31.7154
21.4443 | -35.2403 | -31.7217
21.4504 | -35.2408 | -31.7281
21.4565 | -35.2411 | -31.7344
21.4626 | -35.2415 | -31.7408
21.4687 | -35.242 | -31.7474
21.4748 | -35.2424 | -31.7537
21.4809 | -35.2429 | -31.7601
21.487 | -35.2433 | -31.7667
21.4931 | -35.2439 | -31.7726
21.4992 | -35.2444 | -31.7787
21.5053 | -35.2448 | -31.7851
21.5114 | -35.2452 | -31.7912
21.5175 | -35.2456 | -31.7971
21.5236 | -35.246 | -31.8036
21.5297 | -35.2464 | -31.8096
21.5358 | -35.2469 | -31.8164
21.5419 | -35.2473 | -31.8226
21.548 | -35.2478 | -31.8288
21.5541 | -35.2482 | -31.8348
21.5602 | -35.2487 | -31.8413
21.5664 | -35.2492 | -31.8474
21.5725 | -35.2496 | -31.8539
21.5786 | -35.2501 | -31.86
21.5847 | -35.2506 | -31.8663
21.5908 | -35.251 | -31.8728
21.5969 | -35.2514 | -31.8788
21.603 | -35.2518 | -31.8852
21.6091 | -35.2523 | -31.8915
21.6152 | -35.2527 | -31.898
21.6213 | -35.2531 | -31.9048
21.6274 | -35.2536 | -31.9111
21.6335 | -35.2541 | -31.9169
21.6396 | -35.2545 | -31.9235
21.6457 | -35.255 | -31.93
21.6518 | -35.2553 | -31.9366
21.6579 | -35.2557 | -31.9427
21.664 | -35.2562 | -31.949
21.6701 | -35.2566 | -31.9547
21.6762 | -35.2571 | -31.9605
21.6823 | -35.2576 | -31.9668
21.6884 | -35.258 | -31.9733
21.6945 | -35.2584 | -31.9795
21.7006 | -35.2589 | -31.9862
21.7068 | -35.2593 | -31.9926
21.7129 | -35.2598 | -31.9986
21.719 | -35.2602 | -32.0049
21.7251 | -35.2607 | -32.0115
21.7312 | -35.2611 | -32.0178
21.7373 | -35.2617 | -32.0237
21.7434 | -35.2621 | -32.0303
21.7495 | -35.2626 | -32.0366
21.7556 | -35.263 | -32.0431
21.7617 | -35.2634 | -32.0497
21.7678 | -35.2638 | -32.0561
21.7739 | -35.2643 | -32.0622
21.78 | -35.2647 | -32.0684
21.7861 | -35.2652 | -32.0749
21.7922 | -35.2657 | -32.081
21.7983 | -35.2661 | -32.0871
21.8044 | -35.2666 | -32.0934
21.8105 | -35.267 | -32.0994
21.8166 | -35.2675 | -32.1056
21.8227 | -35.268 | -32.1123
21.8288 | -35.2684 | -32.1185
21.8349 | -35.2688 | -32.1249
21.841 | -35.2692 | -32.131
21.8471 | -35.2696 | -32.1363
21.8533 | -35.2701 | -32.1425
21.8594 | -35.2706 | -32.1486
21.8655 | -35.2711 | -32.1548
21.8716 | -35.2715 | -32.1609
21.8777 | -35.272 | -32.1671
21.8838 | -35.2725 | -32.1732
21.8899 | -35.2729 | -32.1795
21.896 | -35.2734 | -32.1858
21.9021 | -35.2739 | -32.1922
21.9082 | -35.2743 | -32.1988
21.9143 | -35.2747 | -32.2046
21.9204 | -35.2751 | -32.2108
21.9265 | -35.2756 | -32.2169
21.9326 | -35.276 | -32.2229
21.9387 | -35.2764 | -32.2292
21.9448 | -35.2768 | -32.2351
21.9509 | -35.2773 | -32.241
21.957 | -35.2778 | -32.2472
21.9631 | -35.2782 | -32.2534
21.9692 | -35.2786 | -32.2596
21.9753 | -35.2791 | -32.2658
21.9814 | -35.2795 | -32.2718
21.9875 | -35.2799 | -32.2771
21.9937 | -35.2803 | -32.2835
21.9998 | -35.2808 | -32.29
22.0059 | -35.2812 | -32.2963
22.012 | -35.2817 | -32.3026
22.0181 | -35.2822 | -32.3081
22.0242 | -35.2826 | -32.3141
22.0303 | -35.283 | -32.3204
22.0364 | -35.2834 | -32.3264
22.0425 | -35.2838 | -32.3325
22.0486 | -35.2843 | -32.3385
22.0547 | -35.2847 | -32.3447
22.0608 | -35.2852 | -32.3506
22.0669 | -35.2857 | -32.3565
22.073 | -35.2861 | -32.3624
22.0791 | -35.2865 | -32.3688
22.0852 | -35.287 | -32.3751
22.0913 | -35.2873 | -32.3811
22.0974 | -35.2877 | -32.3872
22.1035 | -35.2881 | -32.3934
22.1096 | -35.2885 | -32.3995
22.1157 | -35.289 | -32.4057
22.1218 | -35.2894 | -32.4121
22.1279 | -35.2899 | -32.4178
22.134 | -35.2903 | -32.4242
22.1402 | -35.2908 | -32.4299
22.1463 | -35.2912 | -32.4361
22.1524 | -35.2917 | -32.4424
22.1585 | -35.2921 | -32.4484
22.1646 | -35.2926 | -32.4538
22.1707 | -35.2931 | -32.46
22.1768 | -35.2935 | -32.4662
22.1829 | -35.294 | -32.4724
22.189 | -35.2945 | -32.4787
22.1951 | -35.2949 | -32.4843
22.2012 | -35.2954 | -32.4901
22.2073 | -35.2959 | -32.4961
22.2134 | -35.2963 | -32.502
22.2195 | -35.2968 | -32.5077
22.2256 | -35.2972 | -32.514
22.2317 | -35.2977 | -32.5199
22.2378 | -35.2981 | -32.5261
22.2439 | -35.2986 | -32.5325
22.25 | -35.299 | -32.5385
22.2561 | -35.2994 | -32.5445
22.2622 | -35.2999 | -32.5507
22.2683 | -35.3003 | -32.556
22.2744 | -35.3008 | -32.5621
22.2806 | -35.3012 | -32.5685
22.2867 | -35.3016 | -32.5746
22.2928 | -35.3021 | -32.5808
22.2989 | -35.3025 | -32.5866
22.305 | -35.303 | -32.5925
22.3111 | -35.3034 | -32.5986
22.3172 | -35.3038 | -32.6046
22.3233 | -35.3043 | -32.6106
22.3294 | -35.3048 | -32.6167
22.3355 | -35.3052 | -32.6227
22.3416 | -35.3056 | -32.6289
22.3477 | -35.3061 | -32.635
22.3538 | -35.3065 | -32.641
22.3599 | -35.3069 | -32.6472
22.366 | -35.3074 | -32.653
22.3721 | -35.3078 | -32.6586
22.3782 | -35.3082 | -32.6645
22.3843 | -35.3087 | -32.6708
22.3904 | -35.3091 | -32.6766
22.3965 | -35.3095 | -32.6828
22.4026 | -35.31 | -32.6889
22.4087 | -35.3104 | -32.6946
22.4148 | -35.3108 | -32.7006
22.4209 | -35.3112 | -32.7066
22.4271 | -35.3117 | -32.7124
22.4332 | -35.3122 | -32.7186
22.4393 | -35.3126 | -32.7246
22.4454 | -35.313 | -32.73
22.4515 | -35.3134 | -32.7361
22.4576 | -35.3139 | -32.7423
22.4637 | -35.3144 | -32.7482
22.4698 | -35.3148 | -32.7541
22.4759 | -35.3153 | -32.7602
22.482 | -35.3157 | -32.7662
22.4881 | -35.3161 | -32.7721
22.4942 | -35.3165 | -32.778
22.5003 | -35.317 | -32.7841
22.5064 | -35.3174 | -32.7901
22.5125 | -35.3178 | -32.7955
22.5186 | -35.3183 | -32.8009
22.5247 | -35.3187 | -32.8071
22.5308 | -35.3191 | -32.813
22.5369 | -35.3195 | -32.819
22.543 | -35.3199 | -32.825
22.5491 | -35.3204 | -32.8309
22.5552 | -35.3208 | -32.8367
22.5613 | -35.3212 | -32.8425
22.5675 | -35.3217 | -32.8485
22.5736 | -35.3221 | -32.8548
22.5797 | -35.3226 | -32.8606
22.5858 | -35.323 | -32.8665
22.5919 | -35.3235 | -32.8724
22.598 | -35.324 | -32.8786
22.6041 | -35.3244 | -32.8848
22.6102 | -35.3249 | -32.8908
22.6163 | -35.3253 | -32.8964
22.6224 | -35.3258 | -32.9024
22.6285 | -35.3262 | -32.9085
22.6346 | -35.3267 | -32.9148
22.6407 | -35.3272 | -32.921
22.6468 | -35.3276 | -32.9271
22.6529 | -35.328 | -32.9326
22.659 | -35.3285 | -32.9388
22.6651 | -35.3289 | -32.9447
22.6712 | -35.3293 | -32.9507
22.6773 | -35.3297 | -32.9568
22.6834 | -35.3302 | -32.9626
22.6895 | -35.3306 | -32.9684
22.6956 | -35.331 | -32.9743
22.7017 | -35.3315 | -32.9801
22.7079 | -35.3319 | -32.986
22.714 | -35.3323 | -32.9919
22.7201 | -35.3328 | -32.9978
22.7262 | -35.3332 | -33.0033
22.7323 | -35.3337 | -33.0093
22.7384 | -35.3341 | -33.0154
22.7445 | -35.3345 | -33.0215
22.7506 | -35.335 | -33.0273
22.7567 | -35.3354 | -33.0331
22.7628 | -35.3359 | -33.0388
22.7689 | -35.3363 | -33.0448
22.775 | -35.3367 | -33.0509
22.7811 | -35.3372 | -33.0567
22.7872 | -35.3376 | -33.0626
22.7933 | -35.338 | -33.068
22.7994 | -35.3385 | -33.074
22.8055 | -35.3389 | -33.0798
22.8116 | -35.3393 | -33.0857
22.8177 | -35.3398 | -33.0915
22.8238 | -35.3402 | -33.0976
22.8299 | -35.3407 | -33.1033
22.836 | -35.3411 | -33.1091
22.8421 | -35.3415 | -33.1149
22.8482 | -35.342 | -33.1208
22.8544 | -35.3424 | -33.1266
22.8605 | -35.3428 | -33.1326
22.8666 | -35.3433 | -33.138
22.8727 | -35.3437 | -33.1436
22.8788 | -35.3441 | -33.1493
22.8849 | -35.3445 | -33.155
22.891 | -35.345 | -33.1609
22.8971 | -35.3454 | -33.167
22.9032 | -35.3459 | -33.1724
22.9093 | -35.3463 | -33.1783
22.9154 | -35.3468 | -33.184
22.9215 | -35.3472 | -33.1898
22.9276 | -35.3476 | -33.1957
22.9337 | -35.3481 | -33.2011
22.9398 | -35.3485 | -33.2068
22.9459 | -35.3489 | -33.2128
22.952 | -35.3494 | -33.2186
22.9581 | -35.3499 | -33.2242
22.9642 | -35.3503 | -33.2303
22.9703 | -35.3507 | -33.2358
22.9764 | -35.3511 | -33.2418
22.9825 | -35.3516 | -33.2479
22.9886 | -35.352 | -33.2536
22.9948 | -35.3524 | -33.2593
23.0009 | -35.3529 | -33.2651
23.007 | -35.3534 | -33.2707
23.0131 | -35.3538 | -33.2765
23.0192 | -35.3542 | -33.2825
23.0253 | -35.3546 | -33.2885
23.0314 | -35.355 | -33.2944
23.0375 | -35.3555 | -33.3004
23.0436 | -35.3559 | -33.3061
23.0497 | -35.3563 | -33.312
23.0558 | -35.3568 | -33.3181
23.0619 | -35.3572 | -33.324
23.068 | -35.3576 | -33.3297
23.0741 | -35.3581 | -33.3351
23.0802 | -35.3585 | -33.3412
23.0863 | -35.3589 | -33.3469
23.0924 | -35.3593 | -33.3526
23.0985 | -35.3598 | -33.3585
23.1046 | -35.3602 | -33.3643
23.1107 | -35.3606 | -33.3701
23.1168 | -35.3611 | -33.3757
23.1229 | -35.3615 | -33.3816
23.129 | -35.362 | -33.3872
23.1351 | -35.3624 | -33.3929
23.1413 | -35.3629 | -33.3987
23.1474 | -35.3633 | -33.4043
23.1535 | -35.3637 | -33.4098
23.1596 | -35.3642 | -33.4159
23.1657 | -35.3646 | -33.4217
23.1718 | -35.3651 | -33.4275
23.1779 | -35.3655 | -33.4333
23.184 | -35.3659 | -33.4389
23.1901 | -35.3664 | -33.445
23.1962 | -35.3668 | -33.451
23.2023 | -35.3672 | -33.4567
23.2084 | -35.3677 | -33.4624
23.2145 | -35.3681 | -33.4675
23.2206 | -35.3685 | -33.4735
23.2267 | -35.3689 | -33.4794
23.2328 | -35.3694 | -33.485
23.2389 | -35.3698 | -33.4908
23.245 | -35.3702 | -33.4965
23.2511 | -35.3707 | -33.5021
23.2572 | -35.3711 | -33.5082
23.2633 | -35.3715 | -33.5139
23.2694 | -35.3719 | -33.5199
23.2755 | -35.3724 | -33.5256
23.2817 | -35.3728 | -33.5311
23.2878 | -35.3732 | -33.5369
23.2939 | -35.3736 | -33.5428
23.3 | -35.3741 | -33.5483
23.3061 | -35.3745 | -33.5543
23.3122 | -35.3749 | -33.56
23.3183 | -35.3754 | -33.5659
23.3244 | -35.3758 | -33.5716
23.3305 | -35.3762 | -33.5774
23.3366 | -35.3767 | -33.5831
23.3427 | -35.3771 | -33.5887
23.3488 | -35.3776 | -33.5948
23.3549 | -35.378 | -33.6006
23.361 | -35.3785 | -33.6065
23.3671 | -35.3789 | -33.6123
23.3732 | -35.3793 | -33.6179
23.3793 | -35.3797 | -33.6238
23.3854 | -35.3802 | -33.6295
23.3915 | -35.3806 | -33.6351
23.3976 | -35.3811 | -33.6408
23.4037 | -35.3815 | -33.6465
23.4098 | -35.382 | -33.6522
23.4159 | -35.3824 | -33.6582
23.422 | -35.3829 | -33.664
23.4282 | -35.3833 | -33.6696
23.4343 | -35.3838 | -33.6752
23.4404 | -35.3842 | -33.6807
23.4465 | -35.3846 | -33.6866
23.4526 | -35.385 | -33.6927
23.4587 | -35.3855 | -33.6987
23.4648 | -35.3859 | -33.7043
23.4709 | -35.3863 | -33.71
23.477 | -35.3867 | -33.7156
23.4831 | -35.3871 | -33.7213
23.4892 | -35.3876 | -33.7269
23.4953 | -35.388 | -33.7326
23.5014 | -35.3885 | -33.7382
23.5075 | -35.3889 | -33.744
23.5136 | -35.3893 | -33.7496
23.5197 | -35.3898 | -33.7553
23.5258 | -35.3902 | -33.7609
23.5319 | -35.3906 | -33.7663
23.538 | -35.391 | -33.7722
23.5441 | -35.3915 | -33.778
23.5502 | -35.3919 | -33.7838
23.5563 | -35.3924 | -33.7896
23.5624 | -35.3927 | -33.7954
23.5686 | -35.3932 | -33.8009
23.5747 | -35.3936 | -33.8067
23.5808 | -35.3941 | -33.812
23.5869 | -35.3945 | -33.8178
23.593 | -35.3949 | -33.8235
23.5991 | -35.3954 | -33.8293
23.6052 | -35.3958 | -33.8348
23.6113 | -35.3962 | -33.8408
23.6174 | -35.3966 | -33.8466
23.6235 | -35.397 | -33.8524
23.6296 | -35.3974 | -33.8581
23.6357 | -35.3978 | -33.8638
23.6418 | -35.3982 | -33.8693
23.6479 | -35.3986 | -33.8748
23.654 | -35.3991 | -33.8805
23.6601 | -35.3994 | -33.8866
23.6662 | -35.3999 | -33.8921
23.6723 | -35.4004 | -33.8979
23.6784 | -35.4008 | -33.9036
23.6845 | -35.4012 | -33.9093
23.6906 | -35.4016 | -33.9148
23.6967 | -35.4021 | -33.9203
23.7028 | -35.4026 | -33.9262
23.7089 | -35.403 | -33.932
23.7151 | -35.4034 | -33.9376
23.7212 | -35.4038 | -33.9433
23.7273 | -35.4042 | -33.9492
23.7334 | -35.4047 | -33.9548
23.7395 | -35.4051 | -33.9604
23.7456 | -35.4055 | -33.9659
23.7517 | -35.406 | -33.9718
23.7578 | -35.4064 | -33.9771
23.7639 | -35.4068 | -33.9829
23.77 | -35.4073 | -33.9884
23.7761 | -35.4077 | -33.9943
23.7822 | -35.4081 | -34.0001
23.7883 | -35.4085 | -34.0058
23.7944 | -35.409 | -34.0113
23.8005 | -35.4094 | -34.0172
23.8066 | -35.4098 | -34.0231
23.8127 | -35.4102 | -34.0287
23.8188 | -35.4107 | -34.0342
23.8249 | -35.4111 | -34.0397
23.831 | -35.4115 | -34.0452
23.8371 | -35.412 | -34.0511
23.8432 | -35.4124 | -34.0566
23.8493 | -35.4128 | -34.0619
23.8555 | -35.4133 | -34.0675
23.8616 | -35.4137 | -34.0732
23.8677 | -35.4141 | -34.0786
23.8738 | -35.4145 | -34.0841
23.8799 | -35.415 | -34.0898
23.886 | -35.4154 | -34.0953
23.8921 | -35.4158 | -34.101
23.8982 | -35.4163 | -34.1064
23.9043 | -35.4167 | -34.1121
23.9104 | -35.4171 | -34.1182
23.9165 | -35.4175 | -34.1241
23.9226 | -35.4179 | -34.1297
23.9287 | -35.4183 | -34.1354
23.9348 | -35.4187 | -34.1412
23.9409 | -35.4192 | -34.1466
23.947 | -35.4196 | -34.1523
23.9531 | -35.4201 | -34.158
23.9592 | -35.4204 | -34.1638
23.9653 | -35.4209 | -34.1695
23.9714 | -35.4213 | -34.1754
23.9775 | -35.4217 | -34.1811
23.9836 | -35.4221 | -34.1867
23.9897 | -35.4225 | -34.1924
23.9958 | -35.4229 | -34.1981
24.002 | -35.4233 | -34.2037
24.0081 | -35.4237 | -34.2093
24.0142 | -35.4242 | -34.2147
24.0203 | -35.4246 | -34.2206
24.0264 | -35.425 | -34.2264
24.0325 | -35.4255 | -34.2323
24.0386 | -35.4259 | -34.238
24.0447 | -35.4263 | -34.2439
24.0508 | -35.4268 | -34.2494
24.0569 | -35.4272 | -34.2552
24.063 | -35.4277 | -34.2604
24.0691 | -35.4281 | -34.2663
24.0752 | -35.4285 | -34.2719
24.0813 | -35.429 | -34.2778
24.0874 | -35.4295 | -34.2834
24.0935 | -35.4299 | -34.2891
24.0996 | -35.4303 | -34.2948
24.1057 | -35.4307 | -34.3005
24.1118 | -35.4311 | -34.3064
24.1179 | -35.4315 | -34.3123
24.124 | -35.432 | -34.3182
24.1301 | -35.4324 | -34.3237
24.1362 | -35.4328 | -34.3293
24.1424 | -35.4333 | -34.3348
24.1485 | -35.4337 | -34.3404
24.1546 | -35.4341 | -34.3462
24.1607 | -35.4344 | -34.3519
24.1668 | -35.4349 | -34.3574
24.1729 | -35.4353 | -34.363
24.179 | -35.4357 | -34.3688
24.1851 | -35.4361 | -34.3743
24.1912 | -35.4365 | -34.38
24.1973 | -35.4369 | -34.3858
24.2034 | -35.4374 | -34.3913
24.2095 | -35.4378 | -34.3967
24.2156 | -35.4382 | -34.4023
24.2217 | -35.4387 | -34.4076
24.2278 | -35.4391 | -34.4133
24.2339 | -35.4395 | -34.4189
24.24 | -35.44 | -34.4244
24.2461 | -35.4405 | -34.4298
24.2522 | -35.4408 | -34.4358
24.2583 | -35.4413 | -34.4416
24.2644 | -35.4418 | -34.4476
24.2705 | -35.4422 | -34.4535
24.2766 | -35.4426 | -34.4591
24.2827 | -35.4431 | -34.4647
24.2889 | -35.4435 | -34.4706
24.295 | -35.4439 | -34.4766
24.3011 | -35.4443 | -34.4825
24.3072 | -35.4447 | -34.4881
24.3133 | -35.4452 | -34.4937
24.3194 | -35.4456 | -34.4993
24.3255 | -35.446 | -34.5051
24.3316 | -35.4464 | -34.5108
24.3377 | -35.4468 | -34.5165
24.3438 | -35.4472 | -34.5219
24.3499 | -35.4477 | -34.5277
24.356 | -35.4481 | -34.5334
24.3621 | -35.4485 | -34.5393
24.3682 | -35.4489 | -34.545
24.3743 | -35.4494 | -34.5509
24.3804 | -35.4498 | -34.5564
24.3865 | -35.4502 | -34.5621
24.3926 | -35.4506 | -34.5679
24.3987 | -35.451 | -34.5736
24.4048 | -35.4514 | -34.5793
24.4109 | -35.4518 | -34.5847
24.417 | -35.4522 | -34.5903
24.4231 | -35.4526 | -34.596
24.4293 | -35.453 | -34.6014
24.4354 | -35.4534 | -34.6069
24.4415 | -35.4539 | -34.6126
24.4476 | -35.4543 | -34.6179
24.4537 | -35.4547 | -34.6233
24.4598 | -35.4551 | -34.6292
24.4659 | -35.4556 | -34.6349
24.472 | -35.456 | -34.6405
24.4781 | -35.4565 | -34.6463
24.4842 | -35.4568 | -34.6519
24.4903 | -35.4573 | -34.6576
24.4964 | -35.4576 | -34.663
24.5025 | -35.4581 | -34.6687
24.5086 | -35.4585 | -34.6742
24.5147 | -35.459 | -34.6801
24.5208 | -35.4594 | -34.6854
24.5269 | -35.4598 | -34.6912
24.533 | -35.4602 | -34.6967
24.5391 | -35.4606 | -34.7023
24.5452 | -35.4611 | -34.708
24.5513 | -35.4615 | -34.7139
24.5574 | -35.4619 | -34.7197
24.5635 | -35.4623 | -34.7251
24.5696 | -35.4627 | -34.731
24.5758 | -35.4631 | -34.7365
24.5819 | -35.4635 | -34.742
24.588 | -35.464 | -34.7478
24.5941 | -35.4644 | -34.7532
24.6002 | -35.4648 | -34.7585
24.6063 | -35.4652 | -34.7644
24.6124 | -35.4656 | -34.7699
24.6185 | -35.466 | -34.7758
24.6246 | -35.4664 | -34.781
24.6307 | -35.4668 | -34.7868
24.6368 | -35.4673 | -34.7927
24.6429 | -35.4677 | -34.7982
24.649 | -35.4682 | -34.8041
24.6551 | -35.4686 | -34.8094
24.6612 | -35.469 | -34.8148
24.6673 | -35.4694 | -34.8203
24.6734 | -35.4699 | -34.8264
24.6795 | -35.4703 | -34.8321
24.6856 | -35.4707 | -34.8377
24.6917 | -35.4711 | -34.8433
24.6978 | -35.4715 | -34.8492
24.7039 | -35.4719 | -34.8549
24.71 | -35.4723 | -34.8607
24.7162 | -35.4727 | -34.8665
24.7223 | -35.4731 | -34.8722
24.7284 | -35.4735 | -34.8774
24.7345 | -35.474 | -34.8833
24.7406 | -35.4743 | -34.8889
24.7467 | -35.4747 | -34.8947
24.7528 | -35.4751 | -34.9005
24.7589 | -35.4756 | -34.9061
24.765 | -35.476 | -34.9118
24.7711 | -35.4764 | -34.9176
24.7772 | -35.4768 | -34.9232
24.7833 | -35.4772 | -34.9291
24.7894 | -35.4776 | -34.9347
24.7955 | -35.4781 | -34.9402
24.8016 | -35.4785 | -34.9459
24.8077 | -35.4789 | -34.9515
24.8138 | -35.4793 | -34.9573
24.8199 | -35.4797 | -34.9631
24.826 | -35.4802 | -34.9688
24.8321 | -35.4806 | -34.9742
24.8382 | -35.4809 | -34.9795
24.8443 | -35.4813 | -34.9856
24.8504 | -35.4817 | -34.9913
24.8565 | -35.4821 | -34.9972
24.8627 | -35.4826 | -35.0031
24.8688 | -35.4829 | -35.0087
24.8749 | -35.4833 | -35.0141
24.881 | -35.4838 | -35.0193
24.8871 | -35.4842 | -35.0251
24.8932 | -35.4846 | -35.0306
24.8993 | -35.4851 | -35.036
24.9054 | -35.4855 | -35.0421
24.9115 | -35.4859 | -35.048
24.9176 | -35.4864 | -35.0536
24.9237 | -35.4868 | -35.0594
24.9298 | -35.4871 | -35.065
24.9359 | -35.4876 | -35.0707
24.942 | -35.488 | -35.0761
24.9481 | -35.4884 | -35.0817
24.9542 | -35.4888 | -35.087
24.9603 | -35.4892 | -35.0926
24.9664 | -35.4896 | -35.098
24.9725 | -35.4901 | -35.1037
24.9786 | -35.4905 | -35.1094
24.9847 | -35.4909 | -35.1147
24.9908 | -35.4913 | -35.1208
24.9969 | -35.4917 | -35.1264
25.0031 | -35.4922 | -35.1324
25.0092 | -35.4926 | -35.139
25.0153 | -35.493 | -35.1449
25.0214 | -35.4934 | -35.1509
25.0275 | -35.4939 | -35.1568
25.0336 | -35.4943 | -35.1622
25.0397 | -35.4947 | -35.1678
25.0458 | -35.495 | -35.1732
25.0519 | -35.4955 | -35.1793
25.058 | -35.4958 | -35.1851
25.0641 | -35.4962 | -35.1903
25.0702 | -35.4966 | -35.196
25.0763 | -35.4969 | -35.2017
25.0824 | -35.4973 | -35.2074
25.0885 | -35.4977 | -35.2129
25.0946 | -35.4981 | -35.2185
25.1007 | -35.4985 | -35.224
25.1068 | -35.4988 | -35.2295
25.1129 | -35.4993 | -35.2348
25.119 | -35.4996 | -35.2397
25.1251 | -35.5 | -35.2452
25.1312 | -35.5004 | -35.2508
25.1373 | -35.5008 | -35.2564
25.1435 | -35.5011 | -35.2624
25.1496 | -35.5015 | -35.2681
25.1557 | -35.5019 | -35.2739
25.1618 | -35.5023 | -35.2798
25.1679 | -35.5027 | -35.2855
25.174 | -35.5031 | -35.2909
25.1801 | -35.5036 | -35.297
25.1862 | -35.504 | -35.3028
25.1923 | -35.5044 | -35.3085
25.1984 | -35.5048 | -35.3141
25.2045 | -35.5052 | -35.3197
25.2106 | -35.5056 | -35.3255
25.2167 | -35.506 | -35.3311
25.2228 | -35.5064 | -35.3366
25.2289 | -35.5069 | -35.3418
25.235 | -35.5072 | -35.3476
25.2411 | -35.5077 | -35.353
25.2472 | -35.508 | -35.3588
25.2533 | -35.5084 | -35.3645
25.2594 | -35.5087 | -35.3703
25.2655 | -35.5091 | -35.3761
25.2716 | -35.5095 | -35.3819
25.2777 | -35.5099 | -35.3878
25.2838 | -35.5104 | -35.3931
25.29 | -35.5108 | -35.3984
25.2961 | -35.5112 | -35.4038
25.3022 | -35.5116 | -35.4093
25.3083 | -35.5121 | -35.4149
25.3144 | -35.5125 | -35.4208
25.3205 | -35.5129 | -35.4263
25.3266 | -35.5133 | -35.4319
25.3327 | -35.5138 | -35.4373
25.3388 | -35.5142 | -35.443
25.3449 | -35.5146 | -35.4483
25.351 | -35.515 | -35.4541
25.3571 | -35.5153 | -35.4598
25.3632 | -35.5157 | -35.4651
25.3693 | -35.5161 | -35.4706
25.3754 | -35.5166 | -35.4763
25.3815 | -35.517 | -35.482
25.3876 | -35.5174 | -35.4874
25.3937 | -35.5179 | -35.4931
25.3998 | -35.5182 | -35.4984
25.4059 | -35.5187 | -35.504
25.412 | -35.5191 | -35.5096
25.4181 | -35.5195 | -35.515
25.4242 | -35.5198 | -35.5207
25.4304 | -35.5203 | -35.526
25.4365 | -35.5206 | -35.5315
25.4426 | -35.521 | -35.5373
25.4487 | -35.5214 | -35.5432
25.4548 | -35.5218 | -35.5488
25.4609 | -35.5222 | -35.5543
25.467 | -35.5225 | -35.5602
25.4731 | -35.523 | -35.5659
25.4792 | -35.5234 | -35.5715
25.4853 | -35.5238 | -35.577
25.4914 | -35.5242 | -35.5829
25.4975 | -35.5247 | -35.5887
25.5036 | -35.525 | -35.5939
25.5097 | -35.5254 | -35.5994
25.5158 | -35.5258 | -35.6051
25.5219 | -35.5262 | -35.6107
25.528 | -35.5266 | -35.6159
25.5341 | -35.527 | -35.6215
25.5402 | -35.5274 | -35.6272
25.5463 | -35.5277 | -35.6326
25.5524 | -35.5281 | -35.6384
25.5585 | -35.5285 | -35.6436
25.5646 | -35.5289 | -35.6492
25.5707 | -35.5293 | -35.6548
25.5769 | -35.5296 | -35.6602
25.583 | -35.5299 | -35.6658
25.5891 | -35.5303 | -35.671
25.5952 | -35.5307 | -35.6763
25.6013 | -35.5311 | -35.682
25.6074 | -35.5315 | -35.6873
25.6135 | -35.5319 | -35.6929
25.6196 | -35.5323 | -35.6988
25.6257 | -35.5327 | -35.7043
25.6318 | -35.5331 | -35.7097
25.6379 | -35.5334 | -35.7158
25.644 | -35.5338 | -35.7213
25.6501 | -35.5342 | -35.727
25.6562 | -35.5345 | -35.7327
25.6623 | -35.5349 | -35.7377
25.6684 | -35.5353 | -35.7433
25.6745 | -35.5357 | -35.7494
25.6806 | -35.536 | -35.7548
25.6867 | -35.5364 | -35.7606
25.6928 | -35.5368 | -35.7661
25.6989 | -35.5372 | -35.7718
25.705 | -35.5376 | -35.7777
25.7111 | -35.538 | -35.7835
25.7173 | -35.5384 | -35.7891
25.7234 | -35.5388 | -35.7949
25.7295 | -35.5392 | -35.8007
25.7356 | -35.5396 | -35.8065
25.7417 | -35.54 | -35.8118
25.7478 | -35.5404 | -35.8172
25.7539 | -35.5408 | -35.8225
25.76 | -35.5412 | -35.828
25.7661 | -35.5415 | -35.8334
25.7722 | -35.5419 | -35.8389
25.7783 | -35.5423 | -35.8444
25.7844 | -35.5427 | -35.8498
25.7905 | -35.5431 | -35.8551
25.7966 | -35.5435 | -35.861
25.8027 | -35.5439 | -35.8665
25.8088 | -35.5443 | -35.872
25.8149 | -35.5446 | -35.8771
25.821 | -35.545 | -35.8828
25.8271 | -35.5454 | -35.8886
25.8332 | -35.5458 | -35.8941
25.8393 | -35.5462 | -35.8996
25.8454 | -35.5465 | -35.9056
25.8515 | -35.5468 | -35.9111
25.8576 | -35.5472 | -35.9166
25.8638 | -35.5477 | -35.9223
25.8699 | -35.5481 | -35.9279
25.876 | -35.5485 | -35.9333
25.8821 | -35.5488 | -35.9392
25.8882 | -35.5492 | -35.9444
25.8943 | -35.5497 | -35.9497
25.9004 | -35.5501 | -35.9553
25.9065 | -35.5504 | -35.9611
25.9126 | -35.5508 | -35.9665
25.9187 | -35.5512 | -35.9717
25.9248 | -35.5516 | -35.9775
25.9309 | -35.5519 | -35.983
25.937 | -35.5524 | -35.9885
25.9431 | -35.5528 | -35.9941
25.9492 | -35.553 | -35.9996
25.9553 | -35.5535 | -36.0051
25.9614 | -35.5538 | -36.0109
25.9675 | -35.5542 | -36.0161
25.9736 | -35.5545 | -36.0217
25.9797 | -35.5549 | -36.0271
25.9858 | -35.5553 | -36.0326
25.9919 | -35.5557 | -36.0387
25.998 | -35.556 | -36.0444
26.0042 | -35.5564 | -36.0496
26.0103 | -35.5568 | -36.0552
26.0164 | -35.5572 | -36.0609
26.0225 | -35.5576 | -36.0661
26.0286 | -35.558 | -36.0716
26.0347 | -35.5583 | -36.0767
26.0408 | -35.5587 | -36.0825
26.0469 | -35.5591 | -36.0876
26.053 | -35.5595 | -36.0931
26.0591 | -35.5598 | -36.0985
26.0652 | -35.5599 | -36.1039
26.0713 | -35.5604 | -36.109
26.0774 | -35.5607 | -36.1143
26.0835 | -35.5611 | -36.1197
26.0896 | -35.5616 | -36.1248
26.0957 | -35.5619 | -36.1302
26.1018 | -35.5623 | -36.1353
26.1079 | -35.5627 | -36.1407
26.114 | -35.5631 | -36.1459
26.1201 | -35.5634 | -36.1513
26.1262 | -35.5638 | -36.1575
26.1323 | -35.5642 | -36.1629
26.1384 | -35.5645 | -36.168
26.1445 | -35.5648 | -36.1737
26.1507 | -35.5652 | -36.179
26.1568 | -35.5656 | -36.1841
26.1629 | -35.566 | -36.1897
26.169 | -35.5664 | -36.1955
26.1751 | -35.5667 | -36.2
26.1812 | -35.5671 | -36.205
26.1873 | -35.5675 | -36.2109
26.1934 | -35.5679 | -36.2167
26.1995 | -35.5683 | -36.2219
26.2056 | -35.5687 | -36.2274
26.2117 | -35.5691 | -36.2331
26.2178 | -35.5694 | -36.2381
26.2239 | -35.5698 | -36.2436
26.23 | -35.5701 | -36.2494
26.2361 | -35.5705 | -36.255
26.2422 | -35.5708 | -36.2614
26.2483 | -35.5712 | -36.2668
26.2544 | -35.5715 | -36.2727
26.2605 | -35.5719 | -36.2783
26.2666 | -35.5722 | -36.2836
26.2727 | -35.5726 | -36.2887
26.2788 | -35.5729 | -36.2938
26.2849 | -35.5734 | -36.2992
26.2911 | -35.5737 | -36.3047
26.2972 | -35.5741 | -36.3104
26.3033 | -35.5745 | -36.3161
26.3094 | -35.5748 | -36.3215
26.3155 | -35.5752 | -36.3264
26.3216 | -35.5756 | -36.3317
26.3277 | -35.5759 | -36.3371
26.3338 | -35.5763 | -36.3424
26.3399 | -35.5767 | -36.3479
26.346 | -35.577 | -36.3535
26.3521 | -35.5773 | -36.3591
26.3582 | -35.5777 | -36.3644
26.3643 | -35.5781 | -36.3699
26.3704 | -35.5785 | -36.3754
26.3765 | -35.5789 | -36.3807
26.3826 | -35.5792 | -36.3859
26.3887 | -35.5796 | -36.3914
26.3948 | -35.58 | -36.3967
26.4009 | -35.5803 | -36.4021
26.407 | -35.5806 | -36.4076
26.4131 | -35.5809 | -36.413
26.4192 | -35.5813 | -36.4181
26.4253 | -35.5816 | -36.4236
26.4314 | -35.582 | -36.429
26.4376 | -35.5825 | -36.4344
26.4437 | -35.5828 | -36.44
26.4498 | -35.5832 | -36.4454
26.4559 | -35.5835 | -36.4504
26.462 | -35.5839 | -36.4559
26.4681 | -35.5842 | -36.4617
26.4742 | -35.5846 | -36.4671
26.4803 | -35.585 | -36.4726
26.4864 | -35.5854 | -36.4778
26.4925 | -35.5857 | -36.4832
26.4986 | -35.5861 | -36.4886
26.5047 | -35.5865 | -36.4941
26.5108 | -35.5869 | -36.4992
26.5169 | -35.5872 | -36.5048
26.523 | -35.5875 | -36.5102
26.5291 | -35.5879 | -36.5158
26.5352 | -35.5883 | -36.5212
26.5413 | -35.5887 | -36.5267
26.5474 | -35.589 | -36.532
26.5535 | -35.5893 | -36.5377
26.5596 | -35.5896 | -36.5432
26.5657 | -35.5899 | -36.5486
26.5718 | -35.5903 | -36.5539
26.578 | -35.5907 | -36.5593
26.5841 | -35.591 | -36.5646
26.5902 | -35.5913 | -36.57
26.5963 | -35.5917 | -36.5754
26.6024 | -35.592 | -36.5809
26.6085 | -35.5925 | -36.586
26.6146 | -35.5929 | -36.5917
26.6207 | -35.5933 | -36.5969
26.6268 | -35.5936 | -36.6022
26.6329 | -35.5939 | -36.6078
26.639 | -35.5943 | -36.6132
26.6451 | -35.5946 | -36.6183
26.6512 | -35.595 | -36.6233
26.6573 | -35.5953 | -36.6282
26.6634 | -35.5957 | -36.6338
26.6695 | -35.5961 | -36.6394
26.6756 | -35.5964 | -36.6445
26.6817 | -35.5968 | -36.6501
26.6878 | -35.5971 | -36.6558
26.6939 | -35.5975 | -36.6613
26.7 | -35.5979 | -36.6659
26.7061 | -35.5982 | -36.6715
26.7122 | -35.5986 | -36.6771
26.7183 | -35.599 | -36.6827
26.7245 | -35.5993 | -36.6881
26.7306 | -35.5996 | -36.6932
26.7367 | -35.6001 | -36.6985
26.7428 | -35.6004 | -36.704
26.7489 | -35.6007 | -36.7095
26.755 | -35.6011 | -36.7147
26.7611 | -35.6014 | -36.7201
26.7672 | -35.6018 | -36.7257
26.7733 | -35.6021 | -36.7304
26.7794 | -35.6024 | -36.7362
26.7855 | -35.6028 | -36.7416
26.7916 | -35.6031 | -36.7467
26.7977 | -35.6035 | -36.7515
26.8038 | -35.6038 | -36.7563
26.8099 | -35.6041 | -36.761
26.816 | -35.6044 | -36.7666
26.8221 | -35.6048 | -36.772
26.8282 | -35.6051 | -36.7771
26.8343 | -35.6054 | -36.7829
26.8404 | -35.6058 | -36.7877
26.8465 | -35.6061 | -36.793
26.8526 | -35.6065 | -36.7984
26.8587 | -35.6068 | -36.8039
26.8649 | -35.6071 | -36.8095
26.871 | -35.6075 | -36.8147
26.8771 | -35.6078 | -36.8194
26.8832 | -35.6081 | -36.8247
26.8893 | -35.6085 | -36.8298
26.8954 | -35.6088 | -36.8348
26.9015 | -35.6091 | -36.8398
26.9076 | -35.6095 | -36.8452
26.9137 | -35.6097 | -36.8497
26.9198 | -35.6101 | -36.855
26.9259 | -35.6105 | -36.8605
26.932 | -35.6108 | -36.8658
26.9381 | -35.6112 | -36.8712
26.9442 | -35.6116 | -36.8759
26.9503 | -35.6119 | -36.8815
26.9564 | -35.6122 | -36.8867
26.9625 | -35.6124 | -36.8923
26.9686 | -35.6127 | -36.8976
26.9747 | -35.6131 | -36.9026
26.9808 | -35.6134 | -36.9078
26.9869 | -35.6137 | -36.913
26.993 | -35.6141 | -36.9183
26.9991 | -35.6144 | -36.9234
27.0052 | -35.6147 | -36.9284
27.0114 | -35.6151 | -36.9334
27.0175 | -35.6153 | -36.9391
27.0236 | -35.6157 | -36.9442
27.0297 | -35.6161 | -36.9493
27.0358 | -35.6164 | -36.9546
27.0419 | -35.6168 | -36.9599
27.048 | -35.6171 | -36.965
27.0541 | -35.6174 | -36.9703
27.0602 | -35.6177 | -36.9754
27.0663 | -35.6181 | -36.981
27.0724 | -35.6184 | -36.9866
27.0785 | -35.6187 | -36.9917
27.0846 | -35.619 | -36.9964
27.0907 | -35.6193 | -37.002
27.0968 | -35.6197 | -37.007
27.1029 | -35.62 | -37.0122
27.109 | -35.6204 | -37.0173
27.1151 | -35.6207 | -37.0226
27.1212 | -35.6208 | -37.0276
27.1273 | -35.6212 | -37.0327
27.1334 | -35.6215 | -37.0377
27.1395 | -35.6218 | -37.043
27.1456 | -35.6222 | -37.0483
27.1518 | -35.6225 | -37.0533
27.1579 | -35.6228 | -37.0583
27.164 | -35.6231 | -37.064
27.1701 | -35.6234 | -37.069
27.1762 | -35.6237 | -37.0741
27.1823 | -35.624 | -37.0792
27.1884 | -35.6243 | -37.0839
27.1945 | -35.6246 | -37.0889
27.2006 | -35.6249 | -37.0942
27.2067 | -35.6253 | -37.0994
27.2128 | -35.6256 | -37.1045
27.2189 | -35.6259 | -37.11
27.225 | -35.6262 | -37.115
27.2311 | -35.6264 | -37.1197
27.2372 | -35.6267 | -37.1246
27.2433 | -35.627 | -37.1298
27.2494 | -35.6273 | -37.1351
27.2555 | -35.6276 | -37.1402
27.2616 | -35.6279 | -37.145
27.2677 | -35.6282 | -37.15
27.2738 | -35.6285 | -37.1549
27.2799 | -35.6289 | -37.1598
27.286 | -35.6292 | -37.1649
27.2921 | -35.6296 | -37.1699
27.2983 | -35.6299 | -37.1748
27.3044 | -35.6302 | -37.1803
27.3105 | -35.6306 | -37.1858
27.3166 | -35.6309 | -37.1909
27.3227 | -35.6312 | -37.1958
27.3288 | -35.6315 | -37.201
27.3349 | -35.6318 | -37.2057
27.341 | -35.6322 | -37.2104
27.3471 | -35.6324 | -37.2157
27.3532 | -35.6328 | -37.2206
27.3593 | -35.6332 | -37.2259
27.3654 | -35.6335 | -37.231
27.3715 | -35.6338 | -37.2357
27.3776 | -35.6342 | -37.2409
27.3837 | -35.6345 | -37.246
27.3898 | -35.6348 | -37.251
27.3959 | -35.6351 | -37.2558
27.402 | -35.6354 | -37.2602
27.4081 | -35.6358 | -37.2653
27.4142 | -35.6361 | -37.2702
27.4203 | -35.6364 | -37.2754
27.4264 | -35.6367 | -37.2807
27.4325 | -35.637 | -37.2857
27.4387 | -35.6374 | -37.2906
27.4448 | -35.6378 | -37.2958
27.4509 | -35.6381 | -37.3007
27.457 | -35.6384 | -37.3058
27.4631 | -35.6388 | -37.3108
27.4692 | -35.6391 | -37.3159
27.4753 | -35.6395 | -37.3203
27.4814 | -35.6398 | -37.3255
27.4875 | -35.6401 | -37.3306
27.4936 | -35.6404 | -37.3352
27.4997 | -35.6407 | -37.3403
27.5058 | -35.641 | -37.3453
27.5119 | -35.6414 | -37.3503
27.518 | -35.6417 | -37.3559
27.5241 | -35.642 | -37.3612
27.5302 | -35.6424 | -37.3661
27.5363 | -35.6427 | -37.3714
27.5424 | -35.6429 | -37.3757
27.5485 | -35.6432 | -37.3809
27.5546 | -35.6435 | -37.386
27.5607 | -35.6439 | -37.391
27.5668 | -35.6442 | -37.396
27.5729 | -35.6445 | -37.4009
27.5791 | -35.6449 | -37.4054
27.5852 | -35.6452 | -37.4101
27.5913 | -35.6455 | -37.4154
27.5974 | -35.6458 | -37.4204
27.6035 | -35.6461 | -37.4255
27.6096 | -35.6465 | -37.4305
27.6157 | -35.6468 | -37.4348
27.6218 | -35.6471 | -37.4402
27.6279 | -35.6474 | -37.4451
27.634 | -35.6477 | -37.4505
27.6401 | -35.648 | -37.4558
27.6462 | -35.6483 | -37.4607
27.6523 | -35.6485 | -37.4656
27.6584 | -35.6489 | -37.4709
27.6645 | -35.6492 | -37.4761
27.6706 | -35.6495 | -37.4809
27.6767 | -35.6497 | -37.486
27.6828 | -35.65 | -37.4903
27.6889 | -35.6503 | -37.4954
27.695 | -35.6506 | -37.5003
27.7011 | -35.6509 | -37.5049
27.7072 | -35.6513 | -37.51
27.7133 | -35.6516 | -37.5149
27.7194 | -35.652 | -37.52
27.7256 | -35.6523 | -37.5253
27.7317 | -35.6527 | -37.5302
27.7378 | -35.6531 | -37.5349
27.7439 | -35.6535 | -37.5401
27.75 | -35.6538 | -37.5448
27.7561 | -35.6541 | -37.5496
27.7622 | -35.6544 | -37.5542
27.7683 | -35.6548 | -37.5589
27.7744 | -35.655 | -37.564
27.7805 | -35.6553 | -37.5691
27.7866 | -35.6556 | -37.5742
27.7927 | -35.6559 | -37.5789
27.7988 | -35.6562 | -37.5839
27.8049 | -35.6565 | -37.5882
27.811 | -35.6569 | -37.5932
27.8171 | -35.6572 | -37.5988
27.8232 | -35.6574 | -37.6033
27.8293 | -35.6577 | -37.6082
27.8354 | -35.6581 | -37.613
27.8415 | -35.6584 | -37.6181
27.8476 | -35.6587 | -37.6231
27.8537 | -35.659 | -37.6284
27.8598 | -35.6593 | -37.6328
27.866 | -35.6596 | -37.6378
27.8721 | -35.66 | -37.6428
27.8782 | -35.6603 | -37.6484
27.8843 | -35.6606 | -37.6533
27.8904 | -35.6609 | -37.6582
27.8965 | -35.6612 | -37.663
27.9026 | -35.6615 | -37.6679
27.9087 | -35.6618 | -37.6725
27.9148 | -35.6621 | -37.6776
27.9209 | -35.6623 | -37.6821
27.927 | -35.6626 | -37.687
27.9331 | -35.6628 | -37.692
27.9392 | -35.6631 | -37.6967
27.9453 | -35.6634 | -37.7021
27.9514 | -35.6637 | -37.7068
27.9575 | -35.6639 | -37.7115
27.9636 | -35.6643 | -37.7166
27.9697 | -35.6647 | -37.7209
27.9758 | -35.665 | -37.7258
27.9819 | -35.6653 | -37.7306
27.988 | -35.6656 | -37.7354
27.9941 | -35.6659 | -37.7395
28.0002 | -35.6662 | -37.7438
28.0063 | -35.6665 | -37.749
28.0125 | -35.6667 | -37.7537
28.0186 | -35.6671 | -37.7584
28.0247 | -35.6674 | -37.7636
28.0308 | -35.6677 | -37.7687
28.0369 | -35.6679 | -37.7738
28.043 | -35.6681 | -37.7795
28.0491 | -35.6684 | -37.7841
28.0552 | -35.6687 | -37.7889
28.0613 | -35.669 | -37.7937
28.0674 | -35.6692 | -37.7986
28.0735 | -35.6694 | -37.8035

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4579 | -42.6875
21.4908 | -35.5002 | -34.8022
16.4948 | -34.8795 | -27.4587
12.6602 | -34.5175 | -21.2191
9.7171 | -34.3013 | -16.3419
7.4582 | -34.1527 | -12.3803
5.7244 | -34.0496 | -9.1778
4.3936 | -33.9815 | -6.5834
3.3722 | -33.968 | -4.5446
2.5883 | -33.9682 | -2.9984
1.9866 | -33.9705 | -1.7631
1.5248 | -33.9729 | -0.77397
1.1703 | -33.9932 | -0.1278
1 | -33.9931 | 0.19196
0.89822 | -33.9825 | 0.46222
0.68941 | -33.9692 | 1.0396
0.52919 | -33.9775 | 1.491
0.40612 | -33.9804 | 2.0962
0.31171 | -33.9581 | 2.5525
0.23928 | -33.9652 | 3.2681
0.18358 | -33.9822 | 3.6852
0.14091 | -33.991 | 4.5703
0.10819 | -33.9894 | 5.7303
0.083017 | -34.0317 | 7.2119
0.063705 | -34.0805 | 9.0918
0.048923 | -34.1451 | 11.6018
0.037479 | -34.2388 | 14.8168
0.028801 | -34.4342 | 18.8437
0.022078 | -34.7524 | 24.1454
0.017023 | -35.2418 | 30.4761
0.013018 | -36.0174 | 38.1661
0.010014 | -37.0694 | 46.0682

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.00384 | -14.2353
0.007679 | -31.3147
0.011519 | -37.6818
0.015359 | -40.9231
0.019198 | -43.8739
0.023038 | -45.549
0.026878 | -47.9001
0.030718 | -49.3449
0.034557 | -50.0417
0.038397 | -50.9156
0.042237 | -52.2957
0.046076 | -54.0071
0.049916 | -55.3786
0.053756 | -56.6658
0.057595 | -56.2645
0.061435 | -57.1765
0.065275 | -58.3804
0.069114 | -58.9954
0.072954 | -58.5487
0.076794 | -58.9262
0.080634 | -60.5934
0.084473 | -59.9228
0.088313 | -59.1106
0.092153 | -60.8719
0.095992 | -61.9139
0.099832 | -63.285
0.10367 | -63.3151
0.10751 | -62.5353
0.11135 | -64.4394
0.11519 | -64.8122
0.11903 | -64.0849
0.12287 | -64.3891
0.12671 | -63.8795
0.13055 | -64.4229
0.13439 | -63.8623
0.13823 | -63.5762
0.14207 | -64.6116
0.14591 | -64.9048
0.14975 | -65.7915
0.15359 | -67.1865
0.15743 | -66.4806
0.16127 | -66.0125
0.16511 | -66.926
0.16895 | -67.4053
0.17279 | -67.408
0.17663 | -68.4222
0.18047 | -69.6452
0.1843 | -69.0546
0.18815 | -69.2097
0.19198 | -69.1335
0.19582 | -69.5981
0.19966 | -69.3628
0.2035 | -68.9687
0.20734 | -68.7695
0.21118 | -68.2473
0.21502 | -69.7004
0.21886 | -70.7459
0.2227 | -71.6526
0.22654 | -71.7653
0.23038 | -70.7525
0.23422 | -70.7825
0.23806 | -71.9356
0.2419 | -71.9318
0.24574 | -71.3219
0.24958 | -72.1337
0.25342 | -71.5369
0.25726 | -71.7679
0.2611 | -72.0642
0.26494 | -70.9838
0.26878 | -71.4379
0.27262 | -73.3862
0.27646 | -73.316
0.2803 | -73.0788
0.28414 | -72.2077
0.28798 | -72.5199
0.29182 | -72.2861
0.29566 | -73.7764
0.2995 | -73.9833
0.30334 | -73.8016
0.30717 | -74.1613
0.31101 | -74.2377
0.31485 | -75.2404
0.3187 | -74.3786
0.32253 | -74.0056
0.32637 | -74.0286
0.33021 | -74.3565
0.33405 | -74.1446
0.33789 | -74.3351
0.34173 | -74.3789
0.34557 | -74.5047
0.34941 | -74.6603
0.35325 | -75.2193
0.35709 | -75.8089
0.36093 | -76.5785
0.36477 | -76.3354
0.36861 | -76.2353
0.37245 | -76.7538
0.37629 | -76.4202
0.38013 | -76.8617
0.38397 | -77.3307
0.38781 | -76.996
0.39165 | -76.4385
0.39549 | -76.3509
0.39933 | -76.4203
0.40317 | -76.5637
0.40701 | -77.1705
0.41085 | -77.3092
0.41469 | -76.4143
0.41853 | -77.5646
0.42237 | -78.0598
0.42621 | -77.4656
0.43005 | -78.0658
0.43389 | -77.2077
0.43772 | -76.4288
0.44156 | -76.7982
0.4454 | -77.0516
0.44924 | -77.6603
0.45308 | -78.3919
0.45692 | -78.7713
0.46076 | -77.344
0.4646 | -77.2601
0.46844 | -77.8701
0.47228 | -78.1963
0.47612 | -78.7998
0.47996 | -79.175
0.4838 | -78.4464
0.48764 | -77.5017
0.49148 | -77.7732
0.49532 | -78.2702
0.49916 | -77.5739
0.503 | -77.9528
0.50684 | -78.8228
0.51068 | -79.2171
0.51452 | -80.7719
0.51836 | -79.9539
0.5222 | -79.502
0.52604 | -78.1438
0.52988 | -78.9478
0.53372 | -80.1448
0.53756 | -78.9285
0.5414 | -78.5544
0.54524 | -78.9088
0.54908 | -79.1397
0.55292 | -79.7687
0.55675 | -79.633
0.56059 | -79.3605
0.56444 | -79.4441
0.56827 | -79.6286
0.57211 | -79.7647
0.57595 | -79.9302
0.57979 | -80.0998
0.58363 | -80.1193
0.58747 | -80.3023
0.59131 | -80.9941
0.59515 | -81.3395
0.59899 | -81.2939
0.60283 | -81.3663
0.60667 | -81.2276
0.61051 | -80.8178
0.61435 | -80.7151
0.61819 | -80.4429
0.62203 | -80.0965
0.62587 | -80.5374
0.62971 | -81.2253
0.63355 | -81.7379
0.63739 | -81.5483
0.64123 | -80.7964
0.64507 | -80.7307
0.64891 | -81.3568
0.65275 | -81.731
0.65659 | -81.8789
0.66043 | -81.9417
0.66427 | -82.0571
0.66811 | -81.9807
0.67195 | -81.521
0.67579 | -81.2049
0.67963 | -81.4522
0.68346 | -81.5733
0.68731 | -81.4916
0.69115 | -81.8457
0.69498 | -82.3681
0.69882 | -82.4835
0.70266 | -82.6969
0.7065 | -82.8716
0.71034 | -83.0718
0.71418 | -83.0588
0.71802 | -82.712
0.72186 | -82.5814
0.7257 | -83.0487
0.72954 | -83.3217
0.73338 | -83.1493
0.73722 | -83.0068
0.74106 | -82.6639
0.7449 | -82.3864
0.74874 | -82.6498
0.75258 | -83.1198
0.75642 | -83.2099
0.76026 | -83.3469
0.7641 | -83.7269
0.76794 | -83.6999
0.77178 | -83.7179
0.77562 | -83.8336
0.77946 | -83.4557
0.7833 | -82.9587
0.78714 | -82.8334
0.79098 | -83.1384
0.79482 | -83.7418
0.79866 | -83.9208
0.8025 | -83.5671
0.80634 | -83.5381
0.81017 | -84.0724
0.81402 | -84.5027
0.81785 | -84.0765
0.82169 | -83.511
0.82553 | -83.2965
0.82937 | -83.1941
0.83321 | -83.5306
0.83705 | -84.0171
0.84089 | -84.2637
0.84473 | -84.4333
0.84857 | -84.5714
0.85241 | -84.5996
0.85625 | -84.5693
0.86009 | -84.7638
0.86393 | -84.9948
0.86777 | -85.0609
0.87161 | -85.0356
0.87545 | -84.7068
0.87929 | -84.6263
0.88313 | -84.994
0.88697 | -85.2306
0.89081 | -85.4468
0.89465 | -85.5362
0.89849 | -85.593
0.90233 | -85.7386
0.90617 | -85.8823
0.91001 | -85.4186
0.91385 | -84.9798
0.91769 | -85.2165
0.92153 | -85.4504
0.92537 | -85.4682
0.92921 | -85.6548
0.93305 | -85.6488
0.93688 | -85.2853
0.94073 | -85.1675
0.94456 | -85.0629
0.9484 | -84.8598
0.95224 | -84.8873
0.95608 | -85.0434
0.95992 | -84.9613
0.96376 | -84.7509
0.9676 | -84.725
0.97144 | -85.1075
0.97528 | -85.7891
0.97912 | -86.1292
0.98296 | -86.1672
0.9868 | -86.326
0.99064 | -86.7239
0.99448 | -86.7599
0.99832 | -86.2201
1.0022 | -85.556
1.006 | -85.3672
1.0098 | -85.5192
1.0137 | -85.9032
1.0175 | -86.5233
1.0214 | -86.6166
1.0252 | -86.582
1.029 | -86.6344
1.0329 | -86.3782
1.0367 | -86.4649
1.0406 | -86.315
1.0444 | -85.8152
1.0482 | -85.3798
1.0521 | -85.2434
1.0559 | -85.6831
1.0598 | -86.254
1.0636 | -86.541
1.0674 | -86.741
1.0713 | -86.7819
1.0751 | -87.0534
1.079 | -87.083
1.0828 | -86.8505
1.0866 | -87.0724
1.0905 | -87.1021
1.0943 | -86.8509
1.0982 | -86.8398
1.102 | -86.9466
1.1058 | -87.1101
1.1097 | -87.2101
1.1135 | -87.1389
1.1174 | -86.9846
1.1212 | -86.8908
1.125 | -86.8593
1.1289 | -86.8404
1.1327 | -86.7786
1.1365 | -86.785
1.1404 | -86.9997
1.1442 | -87.2699
1.1481 | -87.3589
1.1519 | -87.371
1.1557 | -87.5554
1.1596 | -87.8485
1.1634 | -87.9554
1.1673 | -87.8469
1.1711 | -87.7895
1.1749 | -87.9467
1.1788 | -88.205
1.1826 | -88.3837
1.1865 | -88.3036
1.1903 | -87.8612
1.1941 | -87.3261
1.198 | -87.1216
1.2018 | -87.3593
1.2057 | -87.7401
1.2095 | -87.8794
1.2133 | -87.7564
1.2172 | -87.5822
1.221 | -87.5302
1.2249 | -87.6594
1.2287 | -87.8972
1.2325 | -88.1464
1.2364 | -88.316
1.2402 | -88.3481
1.2441 | -88.3617
1.2479 | -88.5359
1.2517 | -88.8109
1.2556 | -88.8817
1.2594 | -88.5951
1.2633 | -88.2552
1.2671 | -88.1613
1.2709 | -88.2579
1.2748 | -88.4345
1.2786 | -88.7398
1.2825 | -89.2022
1.2863 | -89.6705
1.2901 | -89.8445
1.294 | -89.6399
1.2978 | -89.3428
1.3017 | -89.1939
1.3055 | -89.2892
1.3093 | -89.5963
1.3132 | -89.964
1.317 | -90.1438
1.3209 | -89.9361
1.3247 | -89.4381
1.3285 | -89.0045
1.3324 | -88.914
1.3362 | -89.1596
1.3401 | -89.456
1.3439 | -89.5064
1.3477 | -89.3054
1.3516 | -89.1418
1.3554 | -89.2357
1.3593 | -89.4731
1.3631 | -89.6387
1.3669 | -89.6198
1.3708 | -89.349
1.3746 | -89.013
1.3784 | -88.9401
1.3823 | -89.1102
1.3861 | -89.2175
1.39 | -89.2044
1.3938 | -89.2614
1.3976 | -89.4198
1.4015 | -89.4879
1.4053 | -89.4261
1.4092 | -89.4412
1.413 | -89.6032
1.4168 | -89.7369
1.4207 | -89.6969
1.4245 | -89.5906
1.4284 | -89.5162
1.4322 | -89.334
1.436 | -89.0705
1.4399 | -88.9527
1.4437 | -89.0807
1.4476 | -89.3664
1.4514 | -89.6218
1.4552 | -89.8033
1.4591 | -90.0801
1.4629 | -90.5285
1.4668 | -90.9526
1.4706 | -91.0607
1.4744 | -90.9515
1.4783 | -90.9442
1.4821 | -90.9474
1.486 | -90.7426
1.4898 | -90.4701
1.4936 | -90.2828
1.4975 | -90.1956
1.5013 | -90.2652
1.5052 | -90.4885
1.509 | -90.6056
1.5128 | -90.3812
1.5167 | -90.109
1.5205 | -90.1117
1.5244 | -90.2065
1.5282 | -90.0532
1.532 | -89.7535
1.5359 | -89.7007
1.5397 | -90.0211
1.5436 | -90.4845
1.5474 | -90.7785
1.5512 | -90.7503
1.5551 | -90.5862
1.5589 | -90.5302
1.5628 | -90.5978
1.5666 | -90.7196
1.5704 | -90.8015
1.5743 | -90.8456
1.5781 | -91.0165
1.582 | -91.2211
1.5858 | -91.1789
1.5896 | -90.8951
1.5935 | -90.6204
1.5973 | -90.499
1.6012 | -90.5216
1.605 | -90.6654
1.6088 | -90.9202
1.6127 | -91.1181
1.6165 | -91.1484
1.6203 | -91.1386
1.6242 | -91.1928
1.628 | -91.2836
1.6319 | -91.3688
1.6357 | -91.3717
1.6395 | -91.3163
1.6434 | -91.3635
1.6472 | -91.5456
1.6511 | -91.5329
1.6549 | -91.1281
1.6587 | -91.0374
1.6626 | -91.0813
1.6664 | -91.2362
1.6703 | -91.3612
1.6741 | -91.4181
1.6779 | -91.3966
1.6818 | -91.2697
1.6856 | -91.064
1.6895 | -90.8575
1.6933 | -90.7782
1.6971 | -90.8942
1.701 | -91.1514
1.7048 | -91.4203
1.7087 | -91.5998
1.7125 | -91.6887
1.7163 | -91.7318
1.7202 | -91.7505
1.724 | -91.7082
1.7279 | -91.5851
1.7317 | -91.4479
1.7355 | -91.3932
1.7394 | -91.4659
1.7432 | -91.5954
1.7471 | -91.6528
1.7509 | -91.602
1.7547 | -91.5011
1.7586 | -91.4043
1.7624 | -91.3405
1.7663 | -91.3543
1.7701 | -91.4498
1.7739 | -91.5844
1.7778 | -91.7593
1.7816 | -92.0072
1.7855 | -92.3001
1.7893 | -92.5248
1.7931 | -92.5486
1.797 | -92.4078
1.8008 | -92.2219
1.8047 | -92.0566
1.8085 | -91.9449
1.8123 | -91.9571
1.8162 | -92.0963
1.82 | -92.2612
1.8239 | -92.3139
1.8277 | -92.212
1.8315 | -92.0115
1.8354 | -91.7838
1.8392 | -91.6151
1.8431 | -91.5257
1.8469 | -91.4882
1.8507 | -91.4991
1.8546 | -91.5775
1.8584 | -91.7275
1.8623 | -91.8725
1.8661 | -91.9537
1.8699 | -92.0106
1.8738 | -92.0921
1.8776 | -92.2246
1.8814 | -92.4029
1.8853 | -92.5938
1.8891 | -92.7394
1.893 | -92.8039
1.8968 | -92.7723
1.9006 | -92.6611
1.9045 | -92.5306
1.9083 | -92.4551
1.9122 | -92.4556
1.916 | -92.4983
1.9198 | -92.5711
1.9237 | -92.6672
1.9275 | -92.7417
1.9314 | -92.7701
1.9352 | -92.7802
1.939 | -92.8246
1.9429 | -92.8584
1.9467 | -92.775
1.9506 | -92.5832
1.9544 | -92.4413
1.9582 | -92.4687
1.9621 | -92.6532
1.9659 | -92.8951
1.9698 | -93.0944
1.9736 | -93.205
1.9774 | -93.2277
1.9813 | -93.2115
1.9851 | -93.203
1.989 | -93.2099
1.9928 | -93.2114
1.9966 | -93.1646
2.0005 | -93.0903
2.0043 | -93.0103
2.0082 | -92.907
2.012 | -92.8041
2.0158 | -92.7566
2.0197 | -92.789
2.0235 | -92.9056
2.0274 | -93.0968
2.0312 | -93.2888
2.035 | -93.336
2.0389 | -93.2457
2.0427 | -93.1891
2.0466 | -93.2768
2.0504 | -93.4854
2.0542 | -93.6942
2.0581 | -93.7695
2.0619 | -93.6279
2.0658 | -93.3229
2.0696 | -93.0517
2.0734 | -92.9607
2.0773 | -93.0424
2.0811 | -93.1549
2.085 | -93.1291
2.0888 | -92.9328
2.0926 | -92.6876
2.0965 | -92.5545
2.1003 | -92.5802
2.1042 | -92.7405
2.108 | -93.0168
2.1118 | -93.34
2.1157 | -93.5744
2.1195 | -93.6165
2.1233 | -93.4946
2.1272 | -93.3152
2.131 | -93.1643
2.1349 | -93.0902
2.1387 | -93.0901
2.1425 | -93.1731
2.1464 | -93.3192
2.1502 | -93.4344
2.1541 | -93.4166
2.1579 | -93.2483
2.1617 | -93.0344
2.1656 | -92.8483
2.1694 | -92.7154
2.1733 | -92.6226
2.1771 | -92.5666
2.1809 | -92.5679
2.1848 | -92.6257
2.1886 | -92.7432
2.1925 | -92.9057
2.1963 | -93.0683
2.2001 | -93.2342
2.204 | -93.4185
2.2078 | -93.6022
2.2117 | -93.6697
2.2155 | -93.5148
2.2193 | -93.4069
2.2232 | -93.3117
2.227 | -93.2585
2.2309 | -93.2558
2.2347 | -93.295
2.2385 | -93.3577
2.2424 | -93.4186
2.2462 | -93.477
2.2501 | -93.5495
2.2539 | -93.6311
2.2577 | -93.7134
2.2616 | -93.7797
2.2654 | -93.8291
2.2693 | -93.8523
2.2731 | -93.8398
2.2769 | -93.7798
2.2808 | -93.6903
2.2846 | -93.6195
2.2885 | -93.6158
2.2923 | -93.7018
2.2961 | -93.8609
2.3 | -94.064
2.3038 | -94.2666
2.3077 | -94.3992
2.3115 | -94.4021
2.3153 | -94.2558
2.3192 | -93.9918
2.323 | -93.6989
2.3269 | -93.4615
2.3307 | -93.3385
2.3345 | -93.3767
2.3384 | -93.5562
2.3422 | -93.7637
2.3461 | -93.9035
2.3499 | -93.9443
2.3537 | -93.8974
2.3576 | -93.8355
2.3614 | -93.8256
2.3653 | -93.897
2.3691 | -94.0337
2.3729 | -94.1826
2.3768 | -94.2839
2.3806 | -94.3196
2.3844 | -94.3217
2.3883 | -94.2885
2.3921 | -94.2373
2.396 | -94.205
2.3998 | -94.2014
2.4036 | -94.2383
2.4075 | -94.3102
2.4113 | -94.415
2.4152 | -94.5061
2.419 | -94.5257
2.4228 | -94.4291
2.4267 | -94.2376
2.4305 | -94.0475
2.4344 | -93.9511
2.4382 | -94.0013
2.442 | -94.1844
2.4459 | -94.4451
2.4497 | -94.7132
2.4536 | -94.9345
2.4574 | -95.0694
2.4612 | -95.1146
2.4651 | -95.1173
2.4689 | -95.1194
2.4728 | -95.1475
2.4766 | -95.1582
2.4804 | -95.0818
2.4843 | -94.8865
2.4881 | -94.6215
2.492 | -94.3767
2.4958 | -94.2327
2.4996 | -94.2091
2.5035 | -94.2728
2.5073 | -94.358
2.5112 | -94.428
2.515 | -94.4837
2.5188 | -94.547
2.5227 | -94.6231
2.5265 | -94.6997
2.5304 | -94.7343
2.5342 | -94.7257
2.538 | -94.6876
2.5419 | -94.6508
2.5457 | -94.6418
2.5496 | -94.6917
2.5534 | -94.7739
2.5572 | -94.8472
2.5611 | -94.8848
2.5649 | -94.8681
2.5688 | -94.8153
2.5726 | -94.7766
2.5764 | -94.775
2.5803 | -94.8291
2.5841 | -94.9234
2.588 | -95.017
2.5918 | -95.0542
2.5956 | -95.0132
2.5995 | -94.8869
2.6033 | -94.7209
2.6072 | -94.5908
2.611 | -94.5412
2.6148 | -94.5756
2.6187 | -94.6724
2.6225 | -94.8135
2.6263 | -94.9856
2.6302 | -95.1605
2.634 | -95.2853
2.6379 | -95.3179
2.6417 | -95.2547
2.6455 | -95.1382
2.6494 | -95.0141
2.6532 | -94.9187
2.6571 | -94.8571
2.6609 | -94.8362
2.6647 | -94.8582
2.6686 | -94.9288
2.6724 | -95.0114
2.6763 | -95.063
2.6801 | -95.0543
2.6839 | -95.028
2.6878 | -95.0241
2.6916 | -95.0424
2.6955 | -95.0381
2.6993 | -94.9879
2.7031 | -94.9326
2.707 | -94.9374
2.7108 | -95.0267
2.7147 | -95.1787
2.7185 | -95.3371
2.7223 | -95.4596
2.7262 | -95.5257
2.73 | -95.5341
2.7339 | -95.5291
2.7377 | -95.5173
2.7415 | -95.4942
2.7454 | -95.4372
2.7492 | -95.3627
2.7531 | -95.3186
2.7569 | -95.331
2.7607 | -95.3825
2.7646 | -95.4132
2.7684 | -95.361
2.7723 | -95.3321
2.7761 | -95.309
2.7799 | -95.3004
2.7838 | -95.3351
2.7876 | -95.402
2.7915 | -95.4937
2.7953 | -95.5956
2.7991 | -95.6831
2.803 | -95.7359
2.8068 | -95.7368
2.8107 | -95.7138
2.8145 | -95.6908
2.8183 | -95.682
2.8222 | -95.7342
2.826 | -95.8511
2.8299 | -96.0087
2.8337 | -96.1534
2.8375 | -96.2415
2.8414 | -96.2284
2.8452 | -96.1112
2.8491 | -95.9394
2.8529 | -95.7749
2.8567 | -95.6667
2.8606 | -95.6142
2.8644 | -95.6137
2.8683 | -95.639
2.8721 | -95.6585
2.8759 | -95.6748
2.8798 | -95.6914
2.8836 | -95.7144
2.8874 | -95.7441
2.8913 | -95.7809
2.8951 | -95.8244
2.899 | -95.8514
2.9028 | -95.8325
2.9066 | -95.7524
2.9105 | -95.6315
2.9143 | -95.4908
2.9182 | -95.3766
2.922 | -95.3184
2.9258 | -95.3267
2.9297 | -95.4061
2.9335 | -95.5373
2.9374 | -95.6763
2.9412 | -95.7864
2.945 | -95.8586
2.9489 | -95.9162
2.9527 | -95.9846
2.9566 | -96.0758
2.9604 | -96.162
2.9642 | -96.2094
2.9681 | -96.1934
2.9719 | -96.121
2.9758 | -96.0163
2.9796 | -95.9149
2.9834 | -95.8488
2.9873 | -95.851
2.9911 | -95.9454
2.995 | -96.1162
2.9988 | -96.3456
3.0026 | -96.5746
3.0065 | -96.726
3.0103 | -96.7831
3.0142 | -96.7615
3.018 | -96.6705
3.0218 | -96.5251
3.0257 | -96.3371
3.0295 | -96.1344
3.0334 | -95.9498
3.0372 | -95.8007
3.041 | -95.7143
3.0449 | -95.7089
3.0487 | -95.7671
3.0526 | -95.8896
3.0564 | -96.05
3.0602 | -96.2043
3.0641 | -96.3294
3.0679 | -96.4154
3.0718 | -96.4706
3.0756 | -96.5099
3.0794 | -96.5346
3.0833 | -96.5658
3.0871 | -96.6081
3.091 | -96.6435
3.0948 | -96.6753
3.0986 | -96.6873
3.1025 | -96.6567
3.1063 | -96.5987
3.1102 | -96.5304
3.114 | -96.4614
3.1178 | -96.4172
3.1217 | -96.4173
3.1255 | -96.4749
3.1294 | -96.5915
3.1332 | -96.739
3.137 | -96.898
3.1409 | -97.0167
3.1447 | -97.03
3.1485 | -96.9318
3.1524 | -96.7483
3.1562 | -96.5308
3.1601 | -96.3552
3.1639 | -96.2849
3.1677 | -96.3232
3.1716 | -96.4332
3.1754 | -96.5698
3.1793 | -96.6842
3.1831 | -96.7532
3.1869 | -96.7798
3.1908 | -96.7772
3.1946 | -96.7682
3.1985 | -96.7505
3.2023 | -96.7225
3.2061 | -96.6934
3.21 | -96.6511
3.2138 | -96.5922
3.2177 | -96.5451
3.2215 | -96.5481
3.2253 | -96.6084
3.2292 | -96.7233
3.233 | -96.8685
3.2369 | -97.0048
3.2407 | -97.1229
3.2445 | -97.2053
3.2484 | -97.2532
3.2522 | -97.2819
3.2561 | -97.2919
3.2599 | -97.2796
3.2637 | -97.2483
3.2676 | -97.1982
3.2714 | -97.1546
3.2753 | -97.1435
3.2791 | -97.1602
3.2829 | -97.1703
3.2868 | -97.1394
3.2906 | -97.0569
3.2945 | -96.9441
3.2983 | -96.8639
3.3021 | -96.8483
3.306 | -96.8982
3.3098 | -97.0035
3.3137 | -97.1314
3.3175 | -97.2593
3.3213 | -97.3418
3.3252 | -97.4347
3.329 | -97.5114
3.3329 | -97.5745
3.3367 | -97.6394
3.3405 | -97.6985
3.3444 | -97.7494
3.3482 | -97.7921
3.3521 | -97.7881
3.3559 | -97.7367
3.3597 | -97.6506
3.3636 | -97.5388
3.3674 | -97.4239
3.3712 | -97.3299
3.3751 | -97.269
3.3789 | -97.245
3.3828 | -97.252
3.3866 | -97.2793
3.3904 | -97.309
3.3943 | -97.3431
3.3981 | -97.3981
3.402 | -97.4703
3.4058 | -97.5645
3.4096 | -97.6821
3.4135 | -97.8035
3.4173 | -97.9056
3.4212 | -97.9849
3.425 | -98.0318
3.4288 | -98.033
3.4327 | -97.9921
3.4365 | -97.9132
3.4404 | -97.8138
3.4442 | -97.7268
3.448 | -97.6745
3.4519 | -97.6586
3.4557 | -97.677
3.4596 | -97.6977
3.4634 | -97.6945
3.4672 | -97.6538
3.4711 | -97.5677
3.4749 | -97.4458
3.4788 | -97.3022
3.4826 | -97.1814
3.4864 | -97.1136
3.4903 | -97.0925
3.4941 | -97.1279
3.498 | -97.2045
3.5018 | -97.2985
3.5056 | -97.3889
3.5095 | -97.4631
3.5133 | -97.5194
3.5172 | -97.5686
3.521 | -97.611
3.5248 | -97.6521
3.5287 | -97.7082
3.5325 | -97.7733
3.5364 | -97.8415
3.5402 | -97.9004
3.544 | -97.9674
3.5479 | -98.0484
3.5517 | -98.1131
3.5556 | -98.1481
3.5594 | -98.1622
3.5632 | -98.1568
3.5671 | -98.1282
3.5709 | -98.0999
3.5748 | -98.066
3.5786 | -98.0238
3.5824 | -97.9694
3.5863 | -97.8984
3.5901 | -97.8288
3.594 | -97.777
3.5978 | -97.758
3.6016 | -97.7766
3.6055 | -97.8274
3.6093 | -97.9072
3.6132 | -97.991
3.617 | -98.0535
3.6208 | -98.0986
3.6247 | -98.129
3.6285 | -98.1193
3.6323 | -98.0859
3.6362 | -98.0331
3.64 | -97.9652
3.6439 | -97.8949
3.6477 | -97.8388
3.6515 | -97.8022
3.6554 | -97.7851
3.6592 | -97.7736
3.6631 | -97.7653
3.6669 | -97.7528
3.6707 | -97.7392
3.6746 | -97.727
3.6784 | -97.7048
3.6823 | -97.6686
3.6861 | -97.6314
3.6899 | -97.579
3.6938 | -97.5297
3.6976 | -97.5077
3.7015 | -97.509
3.7053 | -97.5314
3.7091 | -97.5741
3.713 | -97.6224
3.7168 | -97.6769
3.7207 | -97.748
3.7245 | -97.8271
3.7283 | -97.9077
3.7322 | -97.9784
3.736 | -98.0386
3.7399 | -98.0921
3.7437 | -98.1522
3.7475 | -98.2362
3.7514 | -98.3362
3.7552 | -98.4389
3.7591 | -98.5429
3.7629 | -98.6331
3.7667 | -98.6982
3.7706 | -98.7326
3.7744 | -98.736
3.7783 | -98.6963
3.7821 | -98.6129
3.7859 | -98.5121
3.7898 | -98.4097
3.7936 | -98.3275
3.7975 | -98.2986
3.8013 | -98.3345
3.8051 | -98.4356
3.809 | -98.5806
3.8128 | -98.7293
3.8167 | -98.8424
3.8205 | -98.8856
3.8243 | -98.8659
3.8282 | -98.825
3.832 | -98.795
3.8359 | -98.8035
3.8397 | -98.8726
3.8435 | -98.9723
3.8474 | -99.0962
3.8512 | -99.209
3.8551 | -99.2842
3.8589 | -99.3162
3.8627 | -99.3073
3.8666 | -99.2618
3.8704 | -99.1997
3.8742 | -99.093
3.8781 | -98.9946
3.8819 | -98.8819
3.8858 | -98.7671
3.8896 | -98.6523
3.8934 | -98.5508
3.8973 | -98.4835
3.9011 | -98.4503
3.905 | -98.4571
3.9088 | -98.4915
3.9126 | -98.5562
3.9165 | -98.6305
3.9203 | -98.7085
3.9242 | -98.7573
3.928 | -98.7698
3.9318 | -98.7523
3.9357 | -98.7238
3.9395 | -98.7048
3.9434 | -98.699
3.9472 | -98.7372
3.951 | -98.8123
3.9549 | -98.9146
3.9587 | -99.028
3.9626 | -99.119
3.9664 | -99.1783
3.9702 | -99.203
3.9741 | -99.1929
3.9779 | -99.148
3.9818 | -99.0878
3.9856 | -99.0311
3.9894 | -98.9966
3.9933 | -98.9742
3.9971 | -98.9661
4.001 | -98.9737
4.0048 | -98.9927
4.0086 | -99.0239
4.0125 | -99.0447
4.0163 | -99.0641
4.0202 | -99.094
4.024 | -99.128
4.0278 | -99.1836
4.0317 | -99.2399
4.0355 | -99.2825
4.0394 | -99.2929
4.0432 | -99.2556
4.047 | -99.1629
4.0509 | -99.026
4.0547 | -98.8693
4.0586 | -98.7212
4.0624 | -98.6008
4.0662 | -98.5278
4.0701 | -98.5149
4.0739 | -98.5751
4.0778 | -98.6951
4.0816 | -98.8532
4.0854 | -99.0249
4.0893 | -99.1866
4.0931 | -99.3082
4.097 | -99.3823
4.1008 | -99.4059
4.1046 | -99.3977
4.1085 | -99.3729
4.1123 | -99.3392
4.1162 | -99.2976
4.12 | -99.253
4.1238 | -99.1928
4.1277 | -99.127
4.1315 | -99.0574
4.1353 | -98.9915
4.1392 | -98.9297
4.143 | -98.8855
4.1469 | -98.8545
4.1507 | -98.8291
4.1545 | -98.8175
4.1584 | -98.8199
4.1622 | -98.8414
4.1661 | -98.8916
4.1699 | -98.9643
4.1737 | -99.0555
4.1776 | -99.1547
4.1814 | -99.257
4.1853 | -99.3576
4.1891 | -99.4547
4.1929 | -99.5279
4.1968 | -99.5823
4.2006 | -99.6196
4.2045 | -99.6565
4.2083 | -99.6893
4.2121 | -99.7181
4.216 | -99.7339
4.2198 | -99.7299
4.2237 | -99.7111
4.2275 | -99.6748
4.2313 | -99.6218
4.2352 | -99.5715
4.239 | -99.5201
4.2429 | -99.4838
4.2467 | -99.4516
4.2505 | -99.4198
4.2544 | -99.3936
4.2582 | -99.3775
4.2621 | -99.3765
4.2659 | -99.3845
4.2697 | -99.3858
4.2736 | -99.382
4.2774 | -99.3851
4.2813 | -99.3879
4.2851 | -99.3832
4.2889 | -99.3707
4.2928 | -99.3588
4.2966 | -99.3356
4.3005 | -99.3053
4.3043 | -99.2774
4.3081 | -99.2467
4.312 | -99.2292
4.3158 | -99.2196
4.3197 | -99.2109
4.3235 | -99.2069
4.3273 | -99.1932
4.3312 | -99.1645
4.335 | -99.1282
4.3389 | -99.0895
4.3427 | -99.0306
4.3465 | -98.9754
4.3504 | -98.9276
4.3542 | -98.898
4.3581 | -98.9042
4.3619 | -98.9626
4.3657 | -99.0663
4.3696 | -99.2202
4.3734 | -99.4033
4.3773 | -99.5962
4.3811 | -99.7734
4.3849 | -99.908
4.3888 | -99.9788
4.3926 | -99.9986
4.3964 | -99.9885
4.4003 | -99.9616
4.4041 | -99.9272
4.408 | -99.9191
4.4118 | -99.9376
4.4156 | -99.986
4.4195 | -100.0599
4.4233 | -100.1498
4.4272 | -100.2469
4.431 | -100.3031
4.4348 | -100.3635
4.4387 | -100.3945
4.4425 | -100.4037
4.4464 | -100.3794
4.4502 | -100.3403
4.454 | -100.2853
4.4579 | -100.2286
4.4617 | -100.1822
4.4656 | -100.1366
4.4694 | -100.1032
4.4732 | -100.0865
4.4771 | -100.0658
4.4809 | -100.0469
4.4848 | -100.0233
4.4886 | -99.9881
4.4924 | -99.9461
4.4963 | -99.9021
4.5001 | -99.8605
4.504 | -99.8299
4.5078 | -99.8079
4.5116 | -99.8086
4.5155 | -99.8139
4.5193 | -99.816
4.5232 | -99.8132
4.527 | -99.8131
4.5308 | -99.8146
4.5347 | -99.8189
4.5385 | -99.8407
4.5424 | -99.868
4.5462 | -99.9027
4.55 | -99.9549
4.5539 | -100.0064
4.5577 | -100.0564
4.5616 | -100.1038
4.5654 | -100.1418
4.5692 | -100.1634
4.5731 | -100.159
4.5769 | -100.1261
4.5808 | -100.0783
4.5846 | -100.0115
4.5884 | -99.9499
4.5923 | -99.8935
4.5961 | -99.8528
4.6 | -99.8304
4.6038 | -99.8197
4.6076 | -99.8212
4.6115 | -99.8299
4.6153 | -99.8384
4.6192 | -99.8468
4.623 | -99.833
4.6268 | -99.8114
4.6307 | -99.7831
4.6345 | -99.7438
4.6383 | -99.6968
4.6422 | -99.6532
4.646 | -99.6204
4.6499 | -99.5817
4.6537 | -99.5366
4.6575 | -99.4974
4.6614 | -99.4635
4.6652 | -99.4387
4.6691 | -99.4208
4.6729 | -99.4249
4.6767 | -99.4591
4.6806 | -99.5183
4.6844 | -99.5979
4.6883 | -99.6916
4.6921 | -99.7871
4.6959 | -99.8782
4.6998 | -99.9398
4.7036 | -99.973
4.7075 | -99.9775
4.7113 | -99.9622
4.7151 | -99.9262
4.719 | -99.8841
4.7228 | -99.8446
4.7267 | -99.8076
4.7305 | -99.7757
4.7343 | -99.7478
4.7382 | -99.7328
4.742 | -99.7247
4.7459 | -99.7192
4.7497 | -99.7192
4.7535 | -99.7317
4.7574 | -99.7516
4.7612 | -99.7703
4.7651 | -99.7811
4.7689 | -99.7965
4.7727 | -99.8188
4.7766 | -99.844
4.7804 | -99.8713
4.7843 | -99.9027
4.7881 | -99.9433
4.7919 | -99.9873
4.7958 | -100.0328
4.7996 | -100.0645
4.8035 | -100.0763
4.8073 | -100.0681
4.8111 | -100.0298
4.815 | -99.9768
4.8188 | -99.9173
4.8227 | -99.8575
4.8265 | -99.798
4.8303 | -99.7571
4.8342 | -99.7508
4.838 | -99.7741
4.8419 | -99.8216
4.8457 | -99.9034
4.8495 | -100.0111
4.8534 | -100.143
4.8572 | -100.2881
4.8611 | -100.4392
4.8649 | -100.5854
4.8687 | -100.7116
4.8726 | -100.8071
4.8764 | -100.8663
4.8803 | -100.8844
4.8841 | -100.871
4.8879 | -100.8326
4.8918 | -100.7869
4.8956 | -100.7535
4.8994 | -100.7363
4.9033 | -100.7324
4.9071 | -100.7379
4.911 | -100.753
4.9148 | -100.7756
4.9186 | -100.7869
4.9225 | -100.8129
4.9263 | -100.8502
4.9302 | -100.8876
4.934 | -100.9164
4.9378 | -100.9414
4.9417 | -100.9656
4.9455 | -100.9702
4.9494 | -100.9556
4.9532 | -100.9208
4.957 | -100.8687
4.9609 | -100.8066
4.9647 | -100.7473
4.9686 | -100.6925
4.9724 | -100.6499
4.9762 | -100.6233
4.9801 | -100.6096
4.9839 | -100.6282
4.9878 | -100.6403
4.9916 | -100.6656
4.9954 | -100.7036
4.9993 | -100.7429
5.0031 | -100.7846
5.007 | -100.8343
5.0108 | -100.8729
5.0146 | -100.889
5.0185 | -100.8799
5.0223 | -100.8542
5.0262 | -100.805
5.03 | -100.74
5.0338 | -100.6643
5.0377 | -100.5816
5.0415 | -100.4963
5.0454 | -100.4243
5.0492 | -100.3781
5.053 | -100.3589
5.0569 | -100.3714
5.0607 | -100.42
5.0646 | -100.4969
5.0684 | -100.5907
5.0722 | -100.7005
5.0761 | -100.822
5.0799 | -100.9395
5.0838 | -101.0524
5.0876 | -101.1497
5.0914 | -101.2333
5.0953 | -101.3005
5.0991 | -101.3352
5.103 | -101.3255
5.1068 | -101.2759
5.1106 | -101.1998
5.1145 | -101.0964
5.1183 | -100.9716
5.1222 | -100.835
5.126 | -100.6953
5.1298 | -100.5679
5.1337 | -100.4731
5.1375 | -100.4163
5.1413 | -100.3977
5.1452 | -100.4205
5.149 | -100.4847
5.1529 | -100.5709
5.1567 | -100.6745
5.1605 | -100.7866
5.1644 | -100.887
5.1682 | -100.9672
5.1721 | -101.0201
5.1759 | -101.0372
5.1797 | -101.0322
5.1836 | -101.0198
5.1874 | -100.9903
5.1913 | -100.9413
5.1951 | -100.8865
5.1989 | -100.8377
5.2028 | -100.7907
5.2066 | -100.7435
5.2105 | -100.696
5.2143 | -100.6521
5.2181 | -100.6159
5.222 | -100.5923
5.2258 | -100.5756
5.2297 | -100.5617
5.2335 | -100.5532
5.2373 | -100.5475
5.2412 | -100.5514
5.245 | -100.5681
5.2489 | -100.5965
5.2527 | -100.6387
5.2565 | -100.6939
5.2604 | -100.766
5.2642 | -100.8417
5.2681 | -100.9247
5.2719 | -101.0147
5.2757 | -101.099
5.2796 | -101.1659
5.2834 | -101.2126
5.2873 | -101.2401
5.2911 | -101.2424
5.2949 | -101.2166
5.2988 | -101.1795
5.3026 | -101.1261
5.3065 | -101.0714
5.3103 | -101.0271
5.3141 | -100.9963
5.318 | -100.9787
5.3218 | -100.9723
5.3257 | -100.9787
5.3295 | -100.9959
5.3333 | -101.0232
5.3372 | -101.0376
5.341 | -101.0436
5.3449 | -101.042
5.3487 | -101.0266
5.3525 | -101.0042
5.3564 | -100.9803
5.3602 | -100.9535
5.3641 | -100.9267
5.3679 | -100.9004
5.3717 | -100.8814
5.3756 | -100.8837
5.3794 | -100.9
5.3832 | -100.9313
5.3871 | -100.9722
5.3909 | -101.0196
5.3948 | -101.065
5.3986 | -101.0982
5.4024 | -101.1187
5.4063 | -101.1284
5.4101 | -101.1259
5.414 | -101.1049
5.4178 | -101.0822
5.4216 | -101.0499
5.4255 | -101.0142
5.4293 | -100.9825
5.4332 | -100.9552
5.437 | -100.9382
5.4408 | -100.9417
5.4447 | -100.9647
5.4485 | -101.0016
5.4524 | -101.0541
5.4562 | -101.1103
5.46 | -101.1647
5.4639 | -101.2155
5.4677 | -101.2571
5.4716 | -101.2854
5.4754 | -101.2983
5.4792 | -101.2957
5.4831 | -101.2738
5.4869 | -101.2439
5.4908 | -101.2139
5.4946 | -101.1881
5.4984 | -101.1717
5.5023 | -101.1707
5.5061 | -101.1908
5.51 | -101.2269
5.5138 | -101.2666
5.5176 | -101.312
5.5215 | -101.3563
5.5253 | -101.4037
5.5292 | -101.4545
5.533 | -101.5
5.5368 | -101.5511
5.5407 | -101.6045
5.5445 | -101.6623
5.5484 | -101.7145
5.5522 | -101.7495
5.556 | -101.7769
5.5599 | -101.8012
5.5637 | -101.8174
5.5676 | -101.8188
5.5714 | -101.8074
5.5752 | -101.7909
5.5791 | -101.7758
5.5829 | -101.7589
5.5868 | -101.737
5.5906 | -101.7147
5.5944 | -101.6955
5.5983 | -101.6827
5.6021 | -101.6692
5.606 | -101.6567
5.6098 | -101.6436
5.6136 | -101.6245
5.6175 | -101.6051
5.6213 | -101.5808
5.6251 | -101.5452
5.629 | -101.5043
5.6328 | -101.4602
5.6367 | -101.4121
5.6405 | -101.3626
5.6443 | -101.3109
5.6482 | -101.2672
5.652 | -101.2372
5.6559 | -101.2186
5.6597 | -101.204
5.6635 | -101.1947
5.6674 | -101.1973
5.6712 | -101.2057
5.6751 | -101.2194
5.6789 | -101.2351
5.6827 | -101.2591
5.6866 | -101.2877
5.6904 | -101.3159
5.6943 | -101.343
5.6981 | -101.3645
5.7019 | -101.3885
5.7058 | -101.4076
5.7096 | -101.4221
5.7135 | -101.4307
5.7173 | -101.4317
5.7211 | -101.4323
5.725 | -101.4312
5.7288 | -101.429
5.7327 | -101.417
5.7365 | -101.4011
5.7403 | -101.3899
5.7442 | -101.3815
5.748 | -101.3739
5.7519 | -101.3645
5.7557 | -101.3506
5.7595 | -101.3388
5.7634 | -101.3276
5.7672 | -101.3283
5.7711 | -101.3354
5.7749 | -101.3457
5.7787 | -101.3679
5.7826 | -101.3993
5.7864 | -101.4355
5.7903 | -101.4686
5.7941 | -101.4996
5.7979 | -101.5256
5.8018 | -101.5452
5.8056 | -101.5639
5.8095 | -101.573
5.8133 | -101.5778
5.8171 | -101.5795
5.821 | -101.5876
5.8248 | -101.5929
5.8287 | -101.599
5.8325 | -101.6126
5.8363 | -101.6271
5.8402 | -101.6448
5.844 | -101.6597
5.8479 | -101.6768
5.8517 | -101.694
5.8555 | -101.705
5.8594 | -101.7146
5.8632 | -101.7278
5.8671 | -101.729
5.8709 | -101.7256
5.8747 | -101.7198
5.8786 | -101.7125
5.8824 | -101.707
5.8862 | -101.7034
5.8901 | -101.7067
5.8939 | -101.7188
5.8978 | -101.7422
5.9016 | -101.7754
5.9054 | -101.8064
5.9093 | -101.8267
5.9131 | -101.8449
5.917 | -101.861
5.9208 | -101.8746
5.9246 | -101.8881
5.9285 | -101.8961
5.9323 | -101.9138
5.9362 | -101.9317
5.94 | -101.9454
5.9438 | -101.9524
5.9477 | -101.9453
5.9515 | -101.9124
5.9554 | -101.8708
5.9592 | -101.8147
5.963 | -101.7518
5.9669 | -101.6811
5.9707 | -101.6083
5.9746 | -101.5486
5.9784 | -101.4977
5.9822 | -101.4604
5.9861 | -101.4363
5.9899 | -101.4295
5.9938 | -101.4452
5.9976 | -101.4782
6.0014 | -101.5223
6.0053 | -101.5739
6.0091 | -101.6284
6.013 | -101.6855
6.0168 | -101.745
6.0206 | -101.7835
6.0245 | -101.8072
6.0283 | -101.818
6.0322 | -101.8059
6.036 | -101.7831
6.0398 | -101.7628
6.0437 | -101.7331
6.0475 | -101.6951
6.0514 | -101.6718
6.0552 | -101.664
6.059 | -101.6562
6.0629 | -101.6451
6.0667 | -101.6406
6.0706 | -101.63
6.0744 | -101.6209
6.0782 | -101.6164
6.0821 | -101.6162
6.0859 | -101.626
6.0898 | -101.6765
6.0936 | -101.7184
6.0974 | -101.771
6.1013 | -101.8305
6.1051 | -101.8947
6.109 | -101.9589
6.1128 | -102.0254
6.1166 | -102.0932
6.1205 | -102.1573
6.1243 | -102.2161
6.1281 | -102.2659
6.132 | -102.2969
6.1358 | -102.3167
6.1397 | -102.3345
6.1435 | -102.3319
6.1473 | -102.3166
6.1512 | -102.3006
6.155 | -102.2782
6.1589 | -102.2435
6.1627 | -102.2164
6.1665 | -102.1929
6.1704 | -102.1678
6.1742 | -102.152
6.1781 | -102.1383
6.1819 | -102.1266
6.1857 | -102.1139
6.1896 | -102.1063
6.1934 | -102.1065
6.1973 | -102.1146
6.2011 | -102.1212
6.2049 | -102.1257
6.2088 | -102.1335
6.2126 | -102.1452
6.2165 | -102.1571
6.2203 | -102.1694
6.2241 | -102.1828
6.228 | -102.2
6.2318 | -102.2174
6.2357 | -102.234
6.2395 | -102.2507
6.2433 | -102.27
6.2472 | -102.2914
6.251 | -102.3155
6.2549 | -102.3419
6.2587 | -102.3683
6.2625 | -102.3939
6.2664 | -102.4232
6.2702 | -102.4473
6.2741 | -102.4657
6.2779 | -102.4848
6.2817 | -102.5035
6.2856 | -102.5054
6.2894 | -102.4998
6.2933 | -102.4877
6.2971 | -102.4772
6.3009 | -102.4677
6.3048 | -102.4563
6.3086 | -102.4474
6.3125 | -102.4474
6.3163 | -102.4596
6.3201 | -102.4744
6.324 | -102.4913
6.3278 | -102.5127
6.3317 | -102.5339
6.3355 | -102.5585
6.3393 | -102.5696
6.3432 | -102.5669
6.347 | -102.5572
6.3509 | -102.5367
6.3547 | -102.4925
6.3585 | -102.441
6.3624 | -102.3763
6.3662 | -102.3099
6.37 | -102.2485
6.3739 | -102.1923
6.3777 | -102.1484
6.3816 | -102.1188
6.3854 | -102.0993
6.3892 | -102.0936
6.3931 | -102.0929
6.3969 | -102.1
6.4008 | -102.1127
6.4046 | -102.1178
6.4084 | -102.1215
6.4123 | -102.1162
6.4161 | -102.1116
6.42 | -102.1099
6.4238 | -102.1126
6.4276 | -102.1251
6.4315 | -102.1457
6.4353 | -102.1829
6.4392 | -102.2341
6.443 | -102.2897
6.4468 | -102.3521
6.4507 | -102.4293
6.4545 | -102.5032
6.4584 | -102.5724
6.4622 | -102.6324
6.466 | -102.6802
6.4699 | -102.7144
6.4737 | -102.7382
6.4776 | -102.7625
6.4814 | -102.7788
6.4852 | -102.7882
6.4891 | -102.7944
6.4929 | -102.8067
6.4968 | -102.8152
6.5006 | -102.8266
6.5044 | -102.8327
6.5083 | -102.836
6.5121 | -102.838
6.516 | -102.8366
6.5198 | -102.8313
6.5236 | -102.8238
6.5275 | -102.8152
6.5313 | -102.8029
6.5352 | -102.7817
6.539 | -102.7553
6.5428 | -102.7316
6.5467 | -102.706
6.5505 | -102.6912
6.5544 | -102.6883
6.5582 | -102.6897
6.562 | -102.6945
6.5659 | -102.701
6.5697 | -102.7111
6.5736 | -102.7204
6.5774 | -102.7202
6.5812 | -102.7096
6.5851 | -102.6965
6.5889 | -102.6683
6.5928 | -102.634
6.5966 | -102.5899
6.6004 | -102.5415
6.6043 | -102.4884
6.6081 | -102.4321
6.612 | -102.3808
6.6158 | -102.3382
6.6196 | -102.3026
6.6235 | -102.281
6.6273 | -102.2786
6.6311 | -102.2925
6.635 | -102.3177
6.6388 | -102.3524
6.6427 | -102.3909
6.6465 | -102.4189
6.6503 | -102.4429
6.6542 | -102.46
6.658 | -102.4737
6.6619 | -102.49
6.6657 | -102.5015
6.6695 | -102.504
6.6734 | -102.5007
6.6772 | -102.5101
6.6811 | -102.5217
6.6849 | -102.5284
6.6887 | -102.5369
6.6926 | -102.553
6.6964 | -102.5676
6.7003 | -102.5774
6.7041 | -102.5904
6.7079 | -102.5999
6.7118 | -102.6047
6.7156 | -102.6101
6.7195 | -102.6214
6.7233 | -102.6325
6.7271 | -102.6388
6.731 | -102.6449
6.7348 | -102.6559
6.7387 | -102.6734
6.7425 | -102.6954
6.7463 | -102.729
6.7502 | -102.7669
6.754 | -102.8117
6.7579 | -102.8591
6.7617 | -102.9039
6.7655 | -102.9435
6.7694 | -102.9747
6.7732 | -103.0091
6.7771 | -103.04
6.7809 | -103.0499
6.7847 | -103.0557
6.7886 | -103.055
6.7924 | -103.0492
6.7963 | -103.0367
6.8001 | -103.024
6.8039 | -103.0116
6.8078 | -103.0012
6.8116 | -102.9904
6.8155 | -102.984
6.8193 | -102.9744
6.8231 | -102.9643
6.827 | -102.9572
6.8308 | -102.9573
6.8347 | -102.9641
6.8385 | -102.9661
6.8423 | -102.9637
6.8462 | -102.97
6.85 | -102.9697
6.8539 | -102.9782
6.8577 | -102.9857
6.8615 | -102.9973
6.8654 | -103.0065
6.8692 | -103.0095
6.8731 | -103.0145
6.8769 | -103.0055
6.8807 | -102.9844
6.8846 | -102.9729
6.8884 | -102.9564
6.8922 | -102.9365
6.8961 | -102.9191
6.8999 | -102.9025
6.9038 | -102.8918
6.9076 | -102.8845
6.9114 | -102.8769
6.9153 | -102.8772
6.9191 | -102.8767
6.923 | -102.8749
6.9268 | -102.8725
6.9306 | -102.8688
6.9345 | -102.8614
6.9383 | -102.8549
6.9422 | -102.8518
6.946 | -102.8479
6.9498 | -102.8389
6.9537 | -102.8341
6.9575 | -102.8321
6.9614 | -102.8377
6.9652 | -102.8443
6.969 | -102.8585
6.9729 | -102.882
6.9767 | -102.9121
6.9806 | -102.9513
6.9844 | -102.9906
6.9882 | -103.0193
6.9921 | -103.0439
6.9959 | -103.066
6.9998 | -103.0791
7.0036 | -103.0915
7.0074 | -103.1017
7.0113 | -103.111
7.0151 | -103.12
7.019 | -103.1286
7.0228 | -103.1431
7.0266 | -103.1588
7.0305 | -103.1732
7.0343 | -103.183
7.0382 | -103.187
7.042 | -103.1875
7.0458 | -103.1896
7.0497 | -103.1908
7.0535 | -103.1918
7.0574 | -103.1898
7.0612 | -103.1785
7.065 | -103.1615
7.0689 | -103.1441
7.0727 | -103.1185
7.0766 | -103.0855
7.0804 | -103.05
7.0842 | -103.0158
7.0881 | -102.9863
7.0919 | -102.9656
7.0958 | -102.9527
7.0996 | -102.9479
7.1034 | -102.9525
7.1073 | -102.9624
7.1111 | -102.9814
7.115 | -103.008
7.1188 | -103.0441
7.1226 | -103.0841
7.1265 | -103.1247
7.1303 | -103.1646
7.1341 | -103.1942
7.138 | -103.227
7.1418 | -103.2594
7.1457 | -103.281
7.1495 | -103.2953
7.1533 | -103.3074
7.1572 | -103.319
7.161 | -103.3239
7.1649 | -103.3291
7.1687 | -103.3396
7.1725 | -103.3484
7.1764 | -103.3614
7.1802 | -103.3737
7.1841 | -103.3758
7.1879 | -103.3725
7.1917 | -103.3671
7.1956 | -103.3484
7.1994 | -103.3026
7.2033 | -103.2764
7.2071 | -103.2454
7.2109 | -103.2159
7.2148 | -103.1901
7.2186 | -103.1669
7.2225 | -103.1529
7.2263 | -103.1434
7.2301 | -103.1424
7.234 | -103.1447
7.2378 | -103.1512
7.2417 | -103.1716
7.2455 | -103.1884
7.2493 | -103.2044
7.2532 | -103.2169
7.257 | -103.2235
7.2609 | -103.2209
7.2647 | -103.2122
7.2685 | -103.2008
7.2724 | -103.1785
7.2762 | -103.1495
7.2801 | -103.126
7.2839 | -103.0974
7.2877 | -103.0718
7.2916 | -103.0542
7.2954 | -103.0456
7.2993 | -103.0425
7.3031 | -103.0464
7.3069 | -103.0587
7.3108 | -103.0786
7.3146 | -103.0959
7.3185 | -103.1171
7.3223 | -103.1387
7.3261 | -103.1564
7.33 | -103.1707
7.3338 | -103.1801
7.3377 | -103.1858
7.3415 | -103.1882
7.3453 | -103.189
7.3492 | -103.1852
7.353 | -103.1718
7.3569 | -103.1624
7.3607 | -103.1547
7.3645 | -103.1507
7.3684 | -103.1521
7.3722 | -103.1564
7.3761 | -103.1631
7.3799 | -103.1745
7.3837 | -103.195
7.3876 | -103.2175
7.3914 | -103.2401
7.3952 | -103.2673
7.3991 | -103.2903
7.4029 | -103.3105
7.4068 | -103.323
7.4106 | -103.3305
7.4144 | -103.337
7.4183 | -103.3384
7.4221 | -103.3325
7.426 | -103.3225
7.4298 | -103.3174
7.4336 | -103.3123
7.4375 | -103.3091
7.4413 | -103.3084
7.4452 | -103.3081
7.449 | -103.3107
7.4528 | -103.3124
7.4567 | -103.3153
7.4605 | -103.3242
7.4644 | -103.3367
7.4682 | -103.3421
7.472 | -103.3497
7.4759 | -103.364
7.4797 | -103.3709
7.4836 | -103.3768
7.4874 | -103.3829
7.4912 | -103.3844
7.4951 | -103.3787
7.4989 | -103.3662
7.5028 | -103.3502
7.5066 | -103.3306
7.5104 | -103.309
7.5143 | -103.2866
7.5181 | -103.2653
7.522 | -103.2394
7.5258 | -103.2121
7.5296 | -103.1884
7.5335 | -103.1699
7.5373 | -103.1559
7.5412 | -103.138
7.545 | -103.1232
7.5488 | -103.11
7.5527 | -103.0992
7.5565 | -103.0897
7.5604 | -103.076
7.5642 | -103.065
7.568 | -103.0598
7.5719 | -103.0527
7.5757 | -103.0475
7.5796 | -103.0498
7.5834 | -103.0533
7.5872 | -103.0537
7.5911 | -103.059
7.5949 | -103.0678
7.5988 | -103.0776
7.6026 | -103.0872
7.6064 | -103.0965
7.6103 | -103.1113
7.6141 | -103.1251
7.618 | -103.1413
7.6218 | -103.1578
7.6256 | -103.1732
7.6295 | -103.1867
7.6333 | -103.2064
7.6371 | -103.2267
7.641 | -103.2441
7.6448 | -103.2594
7.6487 | -103.273
7.6525 | -103.2846
7.6563 | -103.2946
7.6602 | -103.3084
7.664 | -103.3228
7.6679 | -103.3367
7.6717 | -103.3583
7.6755 | -103.3753
7.6794 | -103.3893
7.6832 | -103.4063
7.6871 | -103.4181
7.6909 | -103.4332
7.6947 | -103.4481
7.6986 | -103.4546
7.7024 | -103.4603
7.7063 | -103.4521
7.7101 | -103.435
7.7139 | -103.4124
7.7178 | -103.3848
7.7216 | -103.3471
7.7255 | -103.3022
7.7293 | -103.2574
7.7331 | -103.2177
7.737 | -103.1813
7.7408 | -103.1503
7.7447 | -103.1295
7.7485 | -103.12
7.7523 | -103.1622
7.7562 | -103.1718
7.76 | -103.1951
7.7639 | -103.232
7.7677 | -103.2767
7.7715 | -103.3233
7.7754 | -103.3695
7.7792 | -103.4125
7.7831 | -103.4529
7.7869 | -103.4868
7.7907 | -103.5151
7.7946 | -103.5342
7.7984 | -103.5447
7.8023 | -103.5541
7.8061 | -103.5468
7.8099 | -103.5324
7.8138 | -103.52
7.8176 | -103.5017
7.8215 | -103.4825
7.8253 | -103.4676
7.8291 | -103.452
7.833 | -103.4368
7.8368 | -103.432
7.8407 | -103.4295
7.8445 | -103.4257
7.8483 | -103.4256
7.8522 | -103.4324
7.856 | -103.4461
7.8599 | -103.4625
7.8637 | -103.4815
7.8675 | -103.4984
7.8714 | -103.5216
7.8752 | -103.5546
7.8791 | -103.5819
7.8829 | -103.6032
7.8867 | -103.6217
7.8906 | -103.6437
7.8944 | -103.6613
7.8982 | -103.6742
7.9021 | -103.6828
7.9059 | -103.688
7.9098 | -103.6897
7.9136 | -103.6904
7.9174 | -103.6919
7.9213 | -103.6865
7.9251 | -103.6792
7.929 | -103.6797
7.9328 | -103.6831
7.9366 | -103.6874
7.9405 | -103.6946
7.9443 | -103.6988
7.9482 | -103.7048
7.952 | -103.7047
7.9558 | -103.6973
7.9597 | -103.6935
7.9635 | -103.6854
7.9674 | -103.671
7.9712 | -103.6527
7.975 | -103.6298
7.9789 | -103.6008
7.9827 | -103.5705
7.9866 | -103.5455
7.9904 | -103.5204
7.9942 | -103.4981
7.9981 | -103.4867
8.0019 | -103.4802
8.0058 | -103.4858
8.0096 | -103.4963
8.0134 | -103.5138
8.0173 | -103.5336
8.0211 | -103.5569
8.025 | -103.5828
8.0288 | -103.6116
8.0326 | -103.6413
8.0365 | -103.6656
8.0403 | -103.6905
8.0442 | -103.7187
8.048 | -103.7334
8.0518 | -103.7437
8.0557 | -103.7527
8.0595 | -103.7598
8.0634 | -103.7603
8.0672 | -103.7572
8.071 | -103.7568
8.0749 | -103.7506
8.0787 | -103.7346
8.0826 | -103.7205
8.0864 | -103.7028
8.0902 | -103.6806
8.0941 | -103.6542
8.0979 | -103.6276
8.1018 | -103.6001
8.1056 | -103.5671
8.1094 | -103.536
8.1133 | -103.5116
8.1171 | -103.4907
8.121 | -103.4803
8.1248 | -103.4772
8.1286 | -103.4877
8.1325 | -103.5049
8.1363 | -103.536
8.1401 | -103.5723
8.144 | -103.612
8.1478 | -103.6576
8.1517 | -103.7006
8.1555 | -103.7498
8.1593 | -103.7963
8.1632 | -103.8309
8.167 | -103.8569
8.1709 | -103.8752
8.1747 | -103.8826
8.1785 | -103.8793
8.1824 | -103.8665
8.1862 | -103.8455
8.1901 | -103.8156
8.1939 | -103.7806
8.1977 | -103.7465
8.2016 | -103.7123
8.2054 | -103.6808
8.2093 | -103.6572
8.2131 | -103.6416
8.2169 | -103.6293
8.2208 | -103.623
8.2246 | -103.63
8.2285 | -103.6448
8.2323 | -103.6621
8.2361 | -103.6821
8.24 | -103.7072
8.2438 | -103.7384
8.2477 | -103.7708
8.2515 | -103.8073
8.2553 | -103.8443
8.2592 | -103.8805
8.263 | -103.9134
8.2669 | -103.94
8.2707 | -103.9669
8.2745 | -103.9864
8.2784 | -103.9966
8.2822 | -104.0009
8.2861 | -104.0058
8.2899 | -104.0005
8.2937 | -103.9927
8.2976 | -103.9817
8.3014 | -103.9689
8.3053 | -103.9508
8.3091 | -103.9414
8.3129 | -103.929
8.3168 | -103.9159
8.3206 | -103.9076
8.3245 | -103.8976
8.3283 | -103.8925
8.3321 | -103.8855
8.336 | -103.8842
8.3398 | -103.8879
8.3437 | -103.8931
8.3475 | -103.894
8.3513 | -103.8961
8.3552 | -103.9
8.359 | -103.9096
8.3629 | -103.9162
8.3667 | -103.9208
8.3705 | -103.9257
8.3744 | -103.9322
8.3782 | -103.9339
8.382 | -103.934
8.3859 | -103.9362
8.3897 | -103.9383
8.3936 | -103.939
8.3974 | -103.9422
8.4012 | -103.9539
8.4051 | -103.9647
8.4089 | -103.9619
8.4128 | -103.9607
8.4166 | -103.9592
8.4204 | -103.9565
8.4243 | -103.9502
8.4281 | -103.9387
8.432 | -103.9205
8.4358 | -103.9007
8.4396 | -103.8862
8.4435 | -103.871
8.4473 | -103.8496
8.4512 | -103.8343
8.455 | -103.8251
8.4588 | -103.8262
8.4627 | -103.8321
8.4665 | -103.8443
8.4704 | -103.867
8.4742 | -103.9017
8.478 | -103.9401
8.4819 | -103.9861
8.4857 | -104.0394
8.4896 | -104.0968
8.4934 | -104.1501
8.4972 | -104.2063
8.5011 | -104.2615
8.5049 | -104.3143
8.5088 | -104.3629
8.5126 | -104.3988
8.5164 | -104.4272
8.5203 | -104.4456
8.5241 | -104.4558
8.528 | -104.4521
8.5318 | -104.4414
8.5356 | -104.4354
8.5395 | -104.4209
8.5433 | -104.3975
8.5472 | -104.3684
8.551 | -104.3399
8.5548 | -104.319
8.5587 | -104.2966
8.5625 | -104.275
8.5664 | -104.2511
8.5702 | -104.2308
8.574 | -104.2184
8.5779 | -104.2068
8.5817 | -104.1961
8.5856 | -104.1883
8.5894 | -104.181
8.5932 | -104.1751
8.5971 | -104.1689
8.6009 | -104.1612
8.6048 | -104.1527
8.6086 | -104.1433
8.6124 | -104.1306
8.6163 | -104.1196
8.6201 | -104.1048
8.624 | -104.082
8.6278 | -104.0592
8.6316 | -104.0356
8.6355 | -104.0107
8.6393 | -103.9893
8.6431 | -103.9696
8.647 | -103.9534
8.6508 | -103.9428
8.6547 | -103.9422
8.6585 | -103.9484
8.6623 | -103.9555
8.6662 | -103.9719
8.67 | -103.9976
8.6739 | -104.0316
8.6777 | -104.0682
8.6815 | -104.1011
8.6854 | -104.1429
8.6892 | -104.1886
8.6931 | -104.2349
8.6969 | -104.2785
8.7007 | -104.3097
8.7046 | -104.333
8.7084 | -104.3534
8.7123 | -104.3698
8.7161 | -104.3782
8.7199 | -104.3809
8.7238 | -104.3808
8.7276 | -104.3785
8.7315 | -104.3766
8.7353 | -104.3698
8.7391 | -104.36
8.743 | -104.3525
8.7468 | -104.3526
8.7507 | -104.3533
8.7545 | -104.3458
8.7583 | -104.3341
8.7622 | -104.3275
8.766 | -104.3224
8.7699 | -104.311
8.7737 | -104.2939
8.7775 | -104.2654
8.7814 | -104.2336
8.7852 | -104.2026
8.7891 | -104.1699
8.7929 | -104.1296
8.7967 | -104.0867
8.8006 | -104.0457
8.8044 | -104.0069
8.8083 | -103.9738
8.8121 | -103.9395
8.8159 | -103.9111
8.8198 | -103.891
8.8236 | -103.88
8.8275 | -103.8762
8.8313 | -103.8744
8.8351 | -103.8778
8.839 | -103.88
8.8428 | -103.8909
8.8467 | -103.9093
8.8505 | -103.9216
8.8543 | -103.9356
8.8582 | -103.9492
8.862 | -103.9544
8.8659 | -103.9652
8.8697 | -103.9732
8.8735 | -103.9779
8.8774 | -103.9774
8.8812 | -103.9752
8.885 | -103.9743
8.8889 | -103.9722
8.8927 | -103.9659
8.8966 | -103.9608
8.9004 | -103.9611
8.9042 | -103.9638
8.9081 | -103.9723
8.9119 | -103.9827
8.9158 | -103.9991
8.9196 | -104.0226
8.9234 | -104.0466
8.9273 | -104.0691
8.9311 | -104.0987
8.935 | -104.1331
8.9388 | -104.1633
8.9426 | -104.1922
8.9465 | -104.2189
8.9503 | -104.2471
8.9542 | -104.2673
8.958 | -104.2911
8.9618 | -104.317
8.9657 | -104.3342
8.9695 | -104.3517
8.9734 | -104.3681
8.9772 | -104.3717
8.981 | -104.3805
8.9849 | -104.3898
8.9887 | -104.3992
8.9926 | -104.4007
8.9964 | -104.4054
9.0002 | -104.4122
9.0041 | -104.422
9.0079 | -104.4328
9.0118 | -104.4464
9.0156 | -104.4593
9.0194 | -104.4722
9.0233 | -104.4849
9.0271 | -104.4944
9.031 | -104.502
9.0348 | -104.5069
9.0386 | -104.5181
9.0425 | -104.5295
9.0463 | -104.5352
9.0502 | -104.5369
9.054 | -104.539
9.0578 | -104.5419
9.0617 | -104.5462
9.0655 | -104.5465
9.0694 | -104.5468
9.0732 | -104.5559
9.077 | -104.567
9.0809 | -104.5755
9.0847 | -104.5828
9.0886 | -104.5917
9.0924 | -104.5952
9.0962 | -104.5914
9.1001 | -104.5919
9.1039 | -104.595
9.1078 | -104.5915
9.1116 | -104.594
9.1154 | -104.5879
9.1193 | -104.577
9.1231 | -104.5719
9.127 | -104.5712
9.1308 | -104.5672
9.1346 | -104.5636
9.1385 | -104.5613
9.1423 | -104.5542
9.1461 | -104.5466
9.15 | -104.5354
9.1538 | -104.5234
9.1577 | -104.505
9.1615 | -104.4845
9.1653 | -104.457
9.1692 | -104.4282
9.173 | -104.3905
9.1769 | -104.351
9.1807 | -104.3141
9.1845 | -104.2752
9.1884 | -104.24
9.1922 | -104.214
9.1961 | -104.1878
9.1999 | -104.1628
9.2037 | -104.1407
9.2076 | -104.1282
9.2114 | -104.1244
9.2153 | -104.131
9.2191 | -104.1441
9.2229 | -104.1647
9.2268 | -104.1866
9.2306 | -104.2185
9.2345 | -104.2509
9.2383 | -104.2805
9.2421 | -104.3107
9.246 | -104.3373
9.2498 | -104.363
9.2537 | -104.3859
9.2575 | -104.3982
9.2613 | -104.4109
9.2652 | -104.4209
9.269 | -104.425
9.2729 | -104.4239
9.2767 | -104.4226
9.2805 | -104.4146
9.2844 | -104.4024
9.2882 | -104.3842
9.2921 | -104.3663
9.2959 | -104.348
9.2997 | -104.3303
9.3036 | -104.3142
9.3074 | -104.3028
9.3113 | -104.298
9.3151 | -104.3
9.3189 | -104.3112
9.3228 | -104.3254
9.3266 | -104.3461
9.3305 | -104.3778
9.3343 | -104.4142
9.3381 | -104.4559
9.342 | -104.4968
9.3458 | -104.5441
9.3497 | -104.5886
9.3535 | -104.6306
9.3573 | -104.6735
9.3612 | -104.71
9.365 | -104.739
9.3689 | -104.7653
9.3727 | -104.7877
9.3765 | -104.8046
9.3804 | -104.8128
9.3842 | -104.8117
9.388 | -104.8021
9.3919 | -104.7886
9.3957 | -104.7649
9.3996 | -104.7372
9.4034 | -104.7073
9.4072 | -104.6787
9.4111 | -104.6544
9.4149 | -104.6316
9.4188 | -104.6057
9.4226 | -104.5831
9.4264 | -104.5619
9.4303 | -104.5404
9.4341 | -104.5152
9.438 | -104.4941
9.4418 | -104.4711
9.4456 | -104.446
9.4495 | -104.4172
9.4533 | -104.3902
9.4572 | -104.3668
9.461 | -104.3515
9.4648 | -104.3347
9.4687 | -104.3189
9.4725 | -104.3071
9.4764 | -104.3028
9.4802 | -104.3014
9.484 | -104.3026
9.4879 | -104.3135
9.4917 | -104.3303
9.4956 | -104.3489
9.4994 | -104.3753
9.5032 | -104.4023
9.5071 | -104.4282
9.5109 | -104.4555
9.5148 | -104.4824
9.5186 | -104.5122
9.5224 | -104.5457
9.5263 | -104.5755
9.5301 | -104.604
9.534 | -104.636
9.5378 | -104.6636
9.5416 | -104.693
9.5455 | -104.7179
9.5493 | -104.7357
9.5532 | -104.7548
9.557 | -104.7758
9.5608 | -104.7901
9.5647 | -104.7961
9.5685 | -104.7968
9.5724 | -104.7955
9.5762 | -104.7948
9.58 | -104.7837
9.5839 | -104.7691
9.5877 | -104.7501
9.5916 | -104.7346
9.5954 | -104.7178
9.5992 | -104.7
9.6031 | -104.683
9.6069 | -104.6688
9.6108 | -104.6603
9.6146 | -104.652
9.6184 | -104.6396
9.6223 | -104.6332
9.6261 | -104.6293
9.6299 | -104.6264
9.6338 | -104.628
9.6376 | -104.6329
9.6415 | -104.6398
9.6453 | -104.651
9.6491 | -104.6683
9.653 | -104.6879
9.6568 | -104.7097
9.6607 | -104.7283
9.6645 | -104.75
9.6683 | -104.7773
9.6722 | -104.7946
9.676 | -104.806
9.6799 | -104.8231
9.6837 | -104.832
9.6875 | -104.8379
9.6914 | -104.8376
9.6952 | -104.8258
9.6991 | -104.8082
9.7029 | -104.7861
9.7067 | -104.7595
9.7106 | -104.7241
9.7144 | -104.6828
9.7183 | -104.6371
9.7221 | -104.5936
9.7259 | -104.5496
9.7298 | -104.5078
9.7336 | -104.4656
9.7375 | -104.4276
9.7413 | -104.3975
9.7451 | -104.3714
9.749 | -104.3551
9.7528 | -104.3442
9.7567 | -104.3449
9.7605 | -104.354
9.7643 | -104.3659
9.7682 | -104.383
9.772 | -104.4041
9.7759 | -104.4309
9.7797 | -104.4582
9.7835 | -104.4895
9.7874 | -104.5219
9.7912 | -104.5505
9.7951 | -104.5783
9.7989 | -104.607
9.8027 | -104.6345
9.8066 | -104.6518
9.8104 | -104.6633
9.8143 | -104.6808
9.8181 | -104.6955
9.8219 | -104.7035
9.8258 | -104.7116
9.8296 | -104.7144
9.8335 | -104.7182
9.8373 | -104.7161
9.8411 | -104.7112
9.845 | -104.7048
9.8488 | -104.6984
9.8527 | -104.6934
9.8565 | -104.6881
9.8603 | -104.6798
9.8642 | -104.6698
9.868 | -104.6607
9.8719 | -104.6484
9.8757 | -104.6365
9.8795 | -104.6271
9.8834 | -104.616
9.8872 | -104.6089
9.891 | -104.6027
9.8949 | -104.5936
9.8987 | -104.5833
9.9026 | -104.5791
9.9064 | -104.5768
9.9102 | -104.5792
9.9141 | -104.5809
9.9179 | -104.5797
9.9218 | -104.5859
9.9256 | -104.5942
9.9294 | -104.6067
9.9333 | -104.6221
9.9371 | -104.6429
9.941 | -104.6634
9.9448 | -104.6861
9.9486 | -104.7154
9.9525 | -104.7482
9.9563 | -104.7835
9.9602 | -104.8187
9.964 | -104.8557
9.9678 | -104.8961
9.9717 | -104.9317
9.9755 | -104.9713
9.9794 | -105.0066
9.9832 | -105.0468
9.987 | -105.082
9.9909 | -105.1145
9.9947 | -105.144
9.9986 | -105.1693
10.0024 | -105.1926
10.0062 | -105.2093
10.0101 | -105.2244
10.0139 | -105.2378
10.0178 | -105.247
10.0216 | -105.2487
10.0254 | -105.2437
10.0293 | -105.2371
10.0331 | -105.226
10.037 | -105.2164
10.0408 | -105.2085
10.0446 | -105.1976
10.0485 | -105.1841
10.0523 | -105.1671
10.0562 | -105.1466
10.06 | -105.1308
10.0638 | -105.1142
10.0677 | -105.0971
10.0715 | -105.08
10.0754 | -105.0607
10.0792 | -105.044
10.083 | -105.0271
10.0869 | -105.0118
10.0907 | -105.0004
10.0946 | -104.9856
10.0984 | -104.9709
10.1022 | -104.9584
10.1061 | -104.9503
10.1099 | -104.9378
10.1138 | -104.9269
10.1176 | -104.9185
10.1214 | -104.911
10.1253 | -104.9025
10.1291 | -104.8923
10.1329 | -104.8793
10.1368 | -104.8711
10.1406 | -104.863
10.1445 | -104.8511
10.1483 | -104.836
10.1521 | -104.8199
10.156 | -104.8095
10.1598 | -104.7981
10.1637 | -104.7853
10.1675 | -104.7743
10.1713 | -104.7635
10.1752 | -104.7548
10.179 | -104.7522
10.1829 | -104.7522
10.1867 | -104.7515
10.1905 | -104.7539
10.1944 | -104.7627
10.1982 | -104.7756
10.2021 | -104.7924
10.2059 | -104.8123
10.2097 | -104.835
10.2136 | -104.8532
10.2174 | -104.8721
10.2213 | -104.8923
10.2251 | -104.9127
10.2289 | -104.9327
10.2328 | -104.9499
10.2366 | -104.9649
10.2405 | -104.9766
10.2443 | -104.9889
10.2481 | -104.998
10.252 | -105.0033
10.2558 | -105.0092
10.2597 | -105.0164
10.2635 | -105.0203
10.2673 | -105.0206
10.2712 | -105.0178
10.275 | -105.0163
10.2789 | -105.0188
10.2827 | -105.0214
10.2865 | -105.0179
10.2904 | -105.0155
10.2942 | -105.0189
10.2981 | -105.0239
10.3019 | -105.0292
10.3057 | -105.037
10.3096 | -105.0406
10.3134 | -105.0438
10.3173 | -105.0518
10.3211 | -105.0626
10.3249 | -105.0709
10.3288 | -105.0735
10.3326 | -105.0775
10.3365 | -105.0801
10.3403 | -105.0814
10.3441 | -105.0856
10.348 | -105.0879
10.3518 | -105.0901
10.3557 | -105.093
10.3595 | -105.0966
10.3633 | -105.1025
10.3672 | -105.1004
10.371 | -105.0978
10.3749 | -105.0942
10.3787 | -105.0897
10.3825 | -105.0849
10.3864 | -105.0759
10.3902 | -105.0627
10.394 | -105.0489
10.3979 | -105.0356
10.4017 | -105.0214
10.4056 | -105.006
10.4094 | -104.9951
10.4132 | -104.9882
10.4171 | -104.9847
10.4209 | -104.9841
10.4248 | -104.986
10.4286 | -104.9932
10.4324 | -105.0051
10.4363 | -105.0205
10.4401 | -105.0414
10.444 | -105.0669
10.4478 | -105.0966
10.4516 | -105.1292
10.4555 | -105.1613
10.4593 | -105.197
10.4632 | -105.2271
10.467 | -105.2515
10.4708 | -105.2724
10.4747 | -105.2926
10.4785 | -105.3067
10.4824 | -105.3137
10.4862 | -105.318
10.49 | -105.3188
10.4939 | -105.3095
10.4977 | -105.2964
10.5016 | -105.2855
10.5054 | -105.2698
10.5092 | -105.2513
10.5131 | -105.2384
10.5169 | -105.2256
10.5208 | -105.2054
10.5246 | -105.1921
10.5284 | -105.1748
10.5323 | -105.1556
10.5361 | -105.1366
10.54 | -105.1167
10.5438 | -105.1003
10.5476 | -105.0819
10.5515 | -105.0589
10.5553 | -105.0355
10.5592 | -105.0159
10.563 | -104.9967
10.5668 | -104.9743
10.5707 | -104.9464
10.5745 | -104.9223
10.5784 | -104.9026
10.5822 | -104.883
10.586 | -104.8621
10.5899 | -104.8463
10.5937 | -104.831
10.5976 | -104.8165
10.6014 | -104.8051
10.6052 | -104.7936
10.6091 | -104.7869
10.6129 | -104.7875
10.6168 | -104.7965
10.6206 | -104.8094
10.6244 | -104.828
10.6283 | -104.8499
10.6321 | -104.8778
10.6359 | -104.9108
10.6398 | -104.9466
10.6436 | -104.9843
10.6475 | -105.0262
10.6513 | -105.0707
10.6551 | -105.1179
10.659 | -105.1634
10.6628 | -105.2069
10.6667 | -105.2466
10.6705 | -105.2871
10.6743 | -105.3284
10.6782 | -105.3619
10.682 | -105.3891
10.6859 | -105.412
10.6897 | -105.4315
10.6935 | -105.4437
10.6974 | -105.4569
10.7012 | -105.4615
10.7051 | -105.458
10.7089 | -105.4554
10.7127 | -105.4504
10.7166 | -105.436
10.7204 | -105.4202
10.7243 | -105.4093
10.7281 | -105.4031
10.7319 | -105.3899
10.7358 | -105.374
10.7396 | -105.3596
10.7435 | -105.3511
10.7473 | -105.3428
10.7511 | -105.3343
10.755 | -105.3307
10.7588 | -105.3314
10.7627 | -105.335
10.7665 | -105.3401
10.7703 | -105.349
10.7742 | -105.3599
10.778 | -105.3748
10.7819 | -105.3899
10.7857 | -105.4067
10.7895 | -105.4262
10.7934 | -105.4422
10.7972 | -105.4554
10.8011 | -105.4709
10.8049 | -105.4845
10.8087 | -105.4945
10.8126 | -105.5011
10.8164 | -105.5054
10.8203 | -105.507
10.8241 | -105.5113
10.8279 | -105.5137
10.8318 | -105.5162
10.8356 | -105.5191
10.8395 | -105.5266
10.8433 | -105.5338
10.8471 | -105.5395
10.851 | -105.5466
10.8548 | -105.5476
10.8587 | -105.55
10.8625 | -105.5566
10.8663 | -105.5617
10.8702 | -105.5651
10.874 | -105.5645
10.8779 | -105.5595
10.8817 | -105.554
10.8855 | -105.5443
10.8894 | -105.5363
10.8932 | -105.5276
10.897 | -105.5106
10.9009 | -105.4901
10.9047 | -105.4705
10.9086 | -105.453
10.9124 | -105.4316
10.9162 | -105.4093
10.9201 | -105.3864
10.9239 | -105.3612
10.9278 | -105.3324
10.9316 | -105.3002
10.9354 | -105.2706
10.9393 | -105.2406
10.9431 | -105.2132
10.947 | -105.1895
10.9508 | -105.1686
10.9546 | -105.1488
10.9585 | -105.1309
10.9623 | -105.1153
10.9662 | -105.0989
10.97 | -105.0834
10.9738 | -105.0724
10.9777 | -105.0653
10.9815 | -105.0544
10.9854 | -105.0464
10.9892 | -105.0407
10.993 | -105.0405
10.9969 | -105.0354
11.0007 | -105.0329
11.0046 | -105.0335
11.0084 | -105.0344
11.0122 | -105.0378
11.0161 | -105.0421
11.0199 | -105.0449
11.0238 | -105.0483
11.0276 | -105.0525
11.0314 | -105.0609
11.0353 | -105.074
11.0391 | -105.0869
11.043 | -105.1015
11.0468 | -105.117
11.0506 | -105.1329
11.0545 | -105.1553
11.0583 | -105.1803
11.0622 | -105.2029
11.066 | -105.2254
11.0698 | -105.2491
11.0737 | -105.2712
11.0775 | -105.2931
11.0814 | -105.3144
11.0852 | -105.3312
11.089 | -105.3472
11.0929 | -105.3639
11.0967 | -105.3736
11.1006 | -105.3825
11.1044 | -105.3871
11.1082 | -105.3847
11.1121 | -105.3817
11.1159 | -105.3786
11.1197 | -105.3678
11.1236 | -105.3577
11.1274 | -105.3545
11.1313 | -105.3469
11.1351 | -105.3387
11.1389 | -105.329
11.1428 | -105.3244
11.1466 | -105.3218
11.1505 | -105.3189
11.1543 | -105.3143
11.1581 | -105.313
11.162 | -105.3135
11.1658 | -105.3143
11.1697 | -105.317
11.1735 | -105.3213
11.1773 | -105.3228
11.1812 | -105.3262
11.185 | -105.3285
11.1889 | -105.3282
11.1927 | -105.3308
11.1965 | -105.3308
11.2004 | -105.3294
11.2042 | -105.3295
11.2081 | -105.3303
11.2119 | -105.3285
11.2157 | -105.3239
11.2196 | -105.3151
11.2234 | -105.3034
11.2273 | -105.2917
11.2311 | -105.2804
11.2349 | -105.2704
11.2388 | -105.2578
11.2426 | -105.2462
11.2465 | -105.237
11.2503 | -105.2321
11.2541 | -105.2298
11.258 | -105.2311
11.2618 | -105.2348
11.2657 | -105.2417
11.2695 | -105.2475
11.2733 | -105.2538
11.2772 | -105.2669
11.281 | -105.282
11.2849 | -105.3021
11.2887 | -105.3203
11.2925 | -105.337
11.2964 | -105.3574
11.3002 | -105.3812
11.3041 | -105.4029
11.3079 | -105.4213
11.3117 | -105.4368
11.3156 | -105.4487
11.3194 | -105.4587
11.3233 | -105.469
11.3271 | -105.4763
11.3309 | -105.4769
11.3348 | -105.4708
11.3386 | -105.4602
11.3425 | -105.4516
11.3463 | -105.4373
11.3501 | -105.4213
11.354 | -105.4044
11.3578 | -105.3799
11.3617 | -105.3552
11.3655 | -105.3266
11.3693 | -105.2987
11.3732 | -105.2707
11.377 | -105.2455
11.3809 | -105.2213
11.3847 | -105.2002
11.3885 | -105.184
11.3924 | -105.1651
11.3962 | -105.1496
11.4 | -105.1392
11.4039 | -105.1309
11.4077 | -105.1266
11.4116 | -105.1257
11.4154 | -105.1308
11.4192 | -105.1377
11.4231 | -105.1505
11.4269 | -105.1676
11.4308 | -105.1878
11.4346 | -105.2174
11.4384 | -105.2533
11.4423 | -105.2872
11.4461 | -105.3183
11.45 | -105.3486
11.4538 | -105.383
11.4576 | -105.4179
11.4615 | -105.451
11.4653 | -105.4797
11.4692 | -105.5097
11.473 | -105.5372
11.4768 | -105.5648
11.4807 | -105.5923
11.4845 | -105.6151
11.4884 | -105.6312
11.4922 | -105.6429
11.496 | -105.6548
11.4999 | -105.6688
11.5037 | -105.6825
11.5076 | -105.6899
11.5114 | -105.6967
11.5152 | -105.7036
11.5191 | -105.7085
11.5229 | -105.7152
11.5268 | -105.7176
11.5306 | -105.7227
11.5344 | -105.7243
11.5383 | -105.7233
11.5421 | -105.7229
11.546 | -105.7207
11.5498 | -105.7142
11.5536 | -105.7067
11.5575 | -105.7018
11.5613 | -105.6948
11.5652 | -105.6839
11.569 | -105.6756
11.5728 | -105.6697
11.5767 | -105.6583
11.5805 | -105.6425
11.5844 | -105.6297
11.5882 | -105.6122
11.592 | -105.5969
11.5959 | -105.5817
11.5997 | -105.5646
11.6036 | -105.5442
11.6074 | -105.523
11.6112 | -105.5014
11.6151 | -105.4808
11.6189 | -105.4581
11.6227 | -105.4379
11.6266 | -105.419
11.6304 | -105.4168
11.6343 | -105.4052
11.6381 | -105.3922
11.6419 | -105.3802
11.6458 | -105.3719
11.6496 | -105.3666
11.6535 | -105.3623
11.6573 | -105.3623
11.6611 | -105.3623
11.665 | -105.3633
11.6688 | -105.3656
11.6727 | -105.3673
11.6765 | -105.3727
11.6803 | -105.38
11.6842 | -105.3846
11.688 | -105.3905
11.6919 | -105.3976
11.6957 | -105.4042
11.6995 | -105.4105
11.7034 | -105.4139
11.7072 | -105.4189
11.7111 | -105.4256
11.7149 | -105.4318
11.7187 | -105.4376
11.7226 | -105.4377
11.7264 | -105.438
11.7303 | -105.4402
11.7341 | -105.443
11.7379 | -105.4411
11.7418 | -105.4383
11.7456 | -105.4415
11.7495 | -105.4441
11.7533 | -105.4412
11.7571 | -105.4371
11.761 | -105.4313
11.7648 | -105.4249
11.7687 | -105.4184
11.7725 | -105.4136
11.7763 | -105.4094
11.7802 | -105.4102
11.784 | -105.4069
11.7879 | -105.4045
11.7917 | -105.3993
11.7955 | -105.397
11.7994 | -105.3947
11.8032 | -105.3954
11.8071 | -105.3994
11.8109 | -105.4079
11.8147 | -105.4178
11.8186 | -105.4255
11.8224 | -105.4359
11.8263 | -105.4467
11.8301 | -105.454
11.8339 | -105.4658
11.8378 | -105.4788
11.8416 | -105.4918
11.8455 | -105.5055
11.8493 | -105.5172
11.8531 | -105.5276
11.857 | -105.5412
11.8608 | -105.5536
11.8647 | -105.5624
11.8685 | -105.5713
11.8723 | -105.5773
11.8762 | -105.5857
11.88 | -105.5874
11.8838 | -105.5824
11.8877 | -105.5726
11.8915 | -105.5594
11.8954 | -105.549
11.8992 | -105.5319
11.903 | -105.5108
11.9069 | -105.4865
11.9107 | -105.464
11.9146 | -105.4385
11.9184 | -105.409
11.9222 | -105.3776
11.9261 | -105.3482
11.9299 | -105.3204
11.9338 | -105.29
11.9376 | -105.262
11.9414 | -105.2375
11.9453 | -105.2146
11.9491 | -105.1922
11.953 | -105.169
11.9568 | -105.1523
11.9606 | -105.1393
11.9645 | -105.1313
11.9683 | -105.1265
11.9722 | -105.1236
11.976 | -105.1262
11.9798 | -105.1318
11.9837 | -105.1406
11.9875 | -105.1525
11.9914 | -105.1659
11.9952 | -105.1792
11.999 | -105.2
12.0029 | -105.2245
12.0067 | -105.2444
12.0106 | -105.2659
12.0144 | -105.2884
12.0182 | -105.3127
12.0221 | -105.3376
12.0259 | -105.3616
12.0298 | -105.3843
12.0336 | -105.4074
12.0374 | -105.431
12.0413 | -105.4501
12.0451 | -105.4687
12.049 | -105.4851
12.0528 | -105.506
12.0566 | -105.5263
12.0605 | -105.5493
12.0643 | -105.571
12.0682 | -105.5854
12.072 | -105.5994
12.0758 | -105.6092
12.0797 | -105.6182
12.0835 | -105.6199
12.0874 | -105.6188
12.0912 | -105.6192
12.095 | -105.6204
12.0989 | -105.619
12.1027 | -105.6141
12.1066 | -105.608
12.1104 | -105.5975
12.1142 | -105.583
12.1181 | -105.5656
12.1219 | -105.5507
12.1258 | -105.5357
12.1296 | -105.519
12.1334 | -105.5025
12.1373 | -105.4879
12.1411 | -105.4772
12.1449 | -105.4693
12.1488 | -105.4597
12.1526 | -105.4506
12.1565 | -105.4474
12.1603 | -105.4482
12.1641 | -105.4552
12.168 | -105.4654
12.1718 | -105.4751
12.1757 | -105.486
12.1795 | -105.5019
12.1833 | -105.5227
12.1872 | -105.5356
12.191 | -105.5487
12.1949 | -105.563
12.1987 | -105.5787
12.2025 | -105.594
12.2064 | -105.6074
12.2102 | -105.6175
12.2141 | -105.6266
12.2179 | -105.6362
12.2217 | -105.6452
12.2256 | -105.6506
12.2294 | -105.6516
12.2333 | -105.6536
12.2371 | -105.6541
12.2409 | -105.6525
12.2448 | -105.6519
12.2486 | -105.6516
12.2525 | -105.6518
12.2563 | -105.6481
12.2601 | -105.6467
12.264 | -105.6448
12.2678 | -105.643
12.2717 | -105.6417
12.2755 | -105.6416
12.2793 | -105.6436
12.2832 | -105.6505
12.287 | -105.6571
12.2909 | -105.6675
12.2947 | -105.6779
12.2985 | -105.6865
12.3024 | -105.6987
12.3062 | -105.708
12.3101 | -105.7143
12.3139 | -105.7219
12.3177 | -105.7304
12.3216 | -105.7369
12.3254 | -105.7459
12.3293 | -105.7561
12.3331 | -105.7623
12.3369 | -105.7632
12.3408 | -105.7609
12.3446 | -105.7565
12.3485 | -105.7498
12.3523 | -105.7459
12.3561 | -105.7379
12.36 | -105.7243
12.3638 | -105.7097
12.3677 | -105.6975
12.3715 | -105.6816
12.3753 | -105.6681
12.3792 | -105.6553
12.383 | -105.6437
12.3868 | -105.6307
12.3907 | -105.6205
12.3945 | -105.6099
12.3984 | -105.5985
12.4022 | -105.589
12.406 | -105.5835
12.4099 | -105.5811
12.4137 | -105.5811
12.4176 | -105.5826
12.4214 | -105.5841
12.4252 | -105.5874
12.4291 | -105.5967
12.4329 | -105.6076
12.4368 | -105.6202
12.4406 | -105.6347
12.4444 | -105.6493
12.4483 | -105.6618
12.4521 | -105.6711
12.456 | -105.6822
12.4598 | -105.6927
12.4636 | -105.6998
12.4675 | -105.7067
12.4713 | -105.7067
12.4752 | -105.7118
12.479 | -105.7128
12.4828 | -105.7092
12.4867 | -105.7083
12.4905 | -105.7043
12.4944 | -105.6941
12.4982 | -105.6821
12.502 | -105.6738
12.5059 | -105.6641
12.5097 | -105.6532
12.5136 | -105.6433
12.5174 | -105.6385
12.5212 | -105.6358
12.5251 | -105.6359
12.5289 | -105.639
12.5328 | -105.6417
12.5366 | -105.6483
12.5404 | -105.6615
12.5443 | -105.6668
12.5481 | -105.6724
12.552 | -105.6845
12.5558 | -105.6949
12.5596 | -105.7055
12.5635 | -105.7152
12.5673 | -105.7247
12.5712 | -105.7352
12.575 | -105.7468
12.5788 | -105.7592
12.5827 | -105.7657
12.5865 | -105.7693
12.5904 | -105.7703
12.5942 | -105.7684
12.598 | -105.7676
12.6019 | -105.768
12.6057 | -105.7682
12.6096 | -105.7643
12.6134 | -105.7604
12.6172 | -105.7599
12.6211 | -105.7576
12.6249 | -105.7545
12.6288 | -105.7558
12.6326 | -105.7639
12.6364 | -105.7687
12.6403 | -105.7747
12.6441 | -105.7833
12.6479 | -105.7929
12.6518 | -105.8072
12.6556 | -105.8158
12.6595 | -105.8275
12.6633 | -105.8394
12.6671 | -105.8479
12.671 | -105.8607
12.6748 | -105.8705
12.6787 | -105.8775
12.6825 | -105.8866
12.6863 | -105.8942
12.6902 | -105.9023
12.694 | -105.9072
12.6979 | -105.9104
12.7017 | -105.913
12.7055 | -105.9179
12.7094 | -105.9215
12.7132 | -105.9235
12.7171 | -105.9246
12.7209 | -105.9254
12.7247 | -105.935
12.7286 | -105.9425
12.7324 | -105.9479
12.7363 | -105.9531
12.7401 | -105.9633
12.7439 | -105.9689
12.7478 | -105.9775
12.7516 | -105.9855
12.7555 | -105.9938
12.7593 | -106.0025
12.7631 | -106.0116
12.767 | -106.0236
12.7708 | -106.0322
12.7747 | -106.0389
12.7785 | -106.0447
12.7823 | -106.0494
12.7862 | -106.0535
12.79 | -106.0529
12.7939 | -106.0496
12.7977 | -106.0464
12.8015 | -106.0394
12.8054 | -106.0318
12.8092 | -106.0206
12.8131 | -105.999
12.8169 | -105.9722
12.8207 | -105.9486
12.8246 | -105.9236
12.8284 | -105.8951
12.8323 | -105.8644
12.8361 | -105.8342
12.8399 | -105.8023
12.8438 | -105.7697
12.8476 | -105.739
12.8515 | -105.7122
12.8553 | -105.6828
12.8591 | -105.6591
12.863 | -105.6398
12.8668 | -105.6141
12.8706 | -105.5926
12.8745 | -105.5752
12.8783 | -105.5639
12.8822 | -105.5576
12.886 | -105.5575
12.8898 | -105.558
12.8937 | -105.5592
12.8975 | -105.5665
12.9014 | -105.5762
12.9052 | -105.5892
12.909 | -105.6026
12.9129 | -105.6194
12.9167 | -105.6379
12.9206 | -105.6595
12.9244 | -105.6869
12.9282 | -105.7155
12.9321 | -105.7419
12.9359 | -105.769
12.9398 | -105.7989
12.9436 | -105.8275
12.9474 | -105.8542
12.9513 | -105.8805
12.9551 | -105.9027
12.959 | -105.9257
12.9628 | -105.9516
12.9666 | -105.9759
12.9705 | -105.9952
12.9743 | -106.0123
12.9782 | -106.0248
12.982 | -106.0333
12.9858 | -106.0393
12.9897 | -106.0446
12.9935 | -106.049
12.9974 | -106.0542
13.0012 | -106.0592
13.005 | -106.0625
13.0089 | -106.0606
13.0127 | -106.0537
13.0166 | -106.0451
13.0204 | -106.0379
13.0242 | -106.028
13.0281 | -106.0199
13.0319 | -106.009
13.0358 | -105.9986
13.0396 | -105.9869
13.0434 | -105.9733
13.0473 | -105.9579
13.0511 | -105.9436
13.055 | -105.9387
13.0588 | -105.9351
13.0626 | -105.929
13.0665 | -105.92
13.0703 | -105.9089
13.0742 | -105.8977
13.078 | -105.8887
13.0818 | -105.8806
13.0857 | -105.869
13.0895 | -105.8579
13.0934 | -105.8457
13.0972 | -105.8316
13.101 | -105.8184
13.1049 | -105.8051
13.1087 | -105.7909
13.1126 | -105.7771
13.1164 | -105.7622
13.1202 | -105.7525
13.1241 | -105.7469
13.1279 | -105.7421
13.1317 | -105.739
13.1356 | -105.7396
13.1394 | -105.7421
13.1433 | -105.7483
13.1471 | -105.756
13.1509 | -105.764
13.1548 | -105.774
13.1586 | -105.7863
13.1625 | -105.7997
13.1663 | -105.8128
13.1701 | -105.8282
13.174 | -105.8438
13.1778 | -105.8584
13.1817 | -105.8704
13.1855 | -105.8824
13.1893 | -105.8974
13.1932 | -105.9095
13.197 | -105.9228
13.2009 | -105.9371
13.2047 | -105.9493
13.2085 | -105.958
13.2124 | -105.9656
13.2162 | -105.9718
13.2201 | -105.9781
13.2239 | -105.9812
13.2277 | -105.9841
13.2316 | -105.9894
13.2354 | -105.9944
13.2393 | -105.9949
13.2431 | -105.9901
13.2469 | -105.9852
13.2508 | -105.9828
13.2546 | -105.9818
13.2585 | -105.983
13.2623 | -105.9831
13.2661 | -105.9842
13.27 | -105.9883
13.2738 | -105.9913
13.2777 | -105.9912
13.2815 | -105.9905
13.2853 | -105.9947
13.2892 | -105.9991
13.293 | -106.008
13.2969 | -106.0108
13.3007 | -106.0153
13.3045 | -106.0223
13.3084 | -106.0302
13.3122 | -106.036
13.3161 | -106.0401
13.3199 | -106.0468
13.3237 | -106.0524
13.3276 | -106.0573
13.3314 | -106.0644
13.3353 | -106.0705
13.3391 | -106.0763
13.3429 | -106.083
13.3468 | -106.0877
13.3506 | -106.092
13.3545 | -106.1008
13.3583 | -106.1136
13.3621 | -106.122
13.366 | -106.1301
13.3698 | -106.1413
13.3736 | -106.1509
13.3775 | -106.1567
13.3813 | -106.1628
13.3852 | -106.1719
13.389 | -106.179
13.3928 | -106.1872
13.3967 | -106.1943
13.4005 | -106.2012
13.4044 | -106.2072
13.4082 | -106.2125
13.412 | -106.2147
13.4159 | -106.2149
13.4197 | -106.2139
13.4236 | -106.2089
13.4274 | -106.2002
13.4312 | -106.1903
13.4351 | -106.182
13.4389 | -106.1712
13.4428 | -106.1563
13.4466 | -106.1402
13.4504 | -106.1235
13.4543 | -106.1055
13.4581 | -106.0885
13.462 | -106.0719
13.4658 | -106.0574
13.4696 | -106.047
13.4735 | -106.0354
13.4773 | -106.0184
13.4812 | -106.0002
13.485 | -105.9869
13.4888 | -105.974
13.4927 | -105.9656
13.4965 | -105.9588
13.5004 | -105.9503
13.5042 | -105.9416
13.508 | -105.9308
13.5119 | -105.9222
13.5157 | -105.9135
13.5196 | -105.9072
13.5234 | -105.8986
13.5272 | -105.8905
13.5311 | -105.8845
13.5349 | -105.8755
13.5388 | -105.8634
13.5426 | -105.8523
13.5464 | -105.842
13.5503 | -105.8313
13.5541 | -105.8214
13.558 | -105.8142
13.5618 | -105.8059
13.5656 | -105.7955
13.5695 | -105.7882
13.5733 | -105.7873
13.5772 | -105.7872
13.581 | -105.7869
13.5848 | -105.7909
13.5887 | -105.7977
13.5925 | -105.8057
13.5964 | -105.8142
13.6002 | -105.8236
13.604 | -105.8374
13.6079 | -105.8552
13.6117 | -105.8753
13.6156 | -105.8949
13.6194 | -105.916
13.6232 | -105.9384
13.6271 | -105.9593
13.6309 | -105.977
13.6347 | -105.9969
13.6386 | -106.0177
13.6424 | -106.036
13.6463 | -106.0526
13.6501 | -106.0705
13.6539 | -106.0859
13.6578 | -106.099
13.6616 | -106.1073
13.6655 | -106.112
13.6693 | -106.118
13.6731 | -106.1249
13.677 | -106.1296
13.6808 | -106.1297
13.6847 | -106.1286
13.6885 | -106.1252
13.6923 | -106.1211
13.6962 | -106.1151
13.7 | -106.109
13.7039 | -106.1007
13.7077 | -106.0897
13.7115 | -106.0772
13.7154 | -106.0618
13.7192 | -106.0495
13.7231 | -106.0382
13.7269 | -106.0229
13.7307 | -106.0104
13.7346 | -106.0009
13.7384 | -105.9922
13.7423 | -105.9834
13.7461 | -105.9716
13.7499 | -105.963
13.7538 | -105.9528
13.7576 | -105.9455
13.7615 | -105.9387
13.7653 | -105.9357
13.7691 | -105.9325
13.773 | -105.9289
13.7768 | -105.928
13.7807 | -105.9289
13.7845 | -105.9296
13.7883 | -105.9328
13.7922 | -105.9384
13.796 | -105.945
13.7999 | -105.9511
13.8037 | -105.9568
13.8075 | -105.9617
13.8114 | -105.9686
13.8152 | -105.9802
13.8191 | -105.993
13.8229 | -106.0062
13.8267 | -106.0217
13.8306 | -106.0344
13.8344 | -106.0482
13.8383 | -106.0627
13.8421 | -106.0772
13.8459 | -106.0932
13.8498 | -106.1083
13.8536 | -106.1228
13.8575 | -106.1364
13.8613 | -106.1528
13.8651 | -106.1696
13.869 | -106.1813
13.8728 | -106.1929
13.8766 | -106.2057
13.8805 | -106.2181
13.8843 | -106.2326
13.8882 | -106.2459
13.892 | -106.2561
13.8958 | -106.2672
13.8997 | -106.2778
13.9035 | -106.2913
13.9074 | -106.2999
13.9112 | -106.3057
13.915 | -106.3134
13.9189 | -106.3178
13.9227 | -106.3199
13.9266 | -106.32
13.9304 | -106.3187
13.9342 | -106.3157
13.9381 | -106.3092
13.9419 | -106.3014
13.9458 | -106.2974
13.9496 | -106.2915
13.9534 | -106.2873
13.9573 | -106.2772
13.9611 | -106.2642
13.965 | -106.2478
13.9688 | -106.2347
13.9726 | -106.222
13.9765 | -106.2043
13.9803 | -106.1853
13.9842 | -106.1675
13.988 | -106.1517
13.9918 | -106.1359
13.9957 | -106.1182
13.9995 | -106.0982
14.0034 | -106.0796
14.0072 | -106.0639
14.011 | -106.0494
14.0149 | -106.0306
14.0187 | -106.0098
14.0226 | -105.9909
14.0264 | -105.9753
14.0302 | -105.9617
14.0341 | -105.9473
14.0379 | -105.9375
14.0418 | -105.9303
14.0456 | -105.9263
14.0494 | -105.9253
14.0533 | -105.9245
14.0571 | -105.9244
14.061 | -105.9276
14.0648 | -105.9331
14.0686 | -105.9407
14.0725 | -105.9537
14.0763 | -105.9664
14.0802 | -105.978
14.084 | -105.9924
14.0878 | -106.0047
14.0917 | -106.0178
14.0955 | -106.0309
14.0994 | -106.0438
14.1032 | -106.0571
14.107 | -106.0709
14.1109 | -106.086
14.1147 | -106.1004
14.1186 | -106.1109
14.1224 | -106.1188
14.1262 | -106.1251
14.1301 | -106.135
14.1339 | -106.1448
14.1377 | -106.1506
14.1416 | -106.1596
14.1454 | -106.1663
14.1493 | -106.1708
14.1531 | -106.1755
14.1569 | -106.1823
14.1608 | -106.1927
14.1646 | -106.2045
14.1685 | -106.2144
14.1723 | -106.2185
14.1761 | -106.2235
14.18 | -106.2334
14.1838 | -106.2429
14.1877 | -106.2484
14.1915 | -106.251
14.1953 | -106.251
14.1992 | -106.2533
14.203 | -106.2606
14.2069 | -106.264
14.2107 | -106.2664
14.2145 | -106.2705
14.2184 | -106.2777
14.2222 | -106.2815
14.2261 | -106.2827
14.2299 | -106.2787
14.2337 | -106.2765
14.2376 | -106.2786
14.2414 | -106.2763
14.2453 | -106.2704
14.2491 | -106.2641
14.2529 | -106.2555
14.2568 | -106.2455
14.2606 | -106.2397
14.2645 | -106.2339
14.2683 | -106.2255
14.2721 | -106.2133
14.276 | -106.2008
14.2798 | -106.1896
14.2837 | -106.1796
14.2875 | -106.1683
14.2913 | -106.1594
14.2952 | -106.1531
14.299 | -106.1474
14.3029 | -106.143
14.3067 | -106.14
14.3105 | -106.1362
14.3144 | -106.135
14.3182 | -106.1349
14.3221 | -106.1372
14.3259 | -106.1431
14.3297 | -106.1472
14.3336 | -106.1515
14.3374 | -106.1552
14.3413 | -106.158
14.3451 | -106.163
14.3489 | -106.1695
14.3528 | -106.1748
14.3566 | -106.1774
14.3605 | -106.1785
14.3643 | -106.184
14.3681 | -106.1888
14.372 | -106.1944
14.3758 | -106.2013
14.3796 | -106.2082
14.3835 | -106.211
14.3873 | -106.2133
14.3912 | -106.2192
14.395 | -106.2268
14.3988 | -106.2305
14.4027 | -106.2391
14.4065 | -106.2447
14.4104 | -106.2486
14.4142 | -106.2546
14.418 | -106.2591
14.4219 | -106.2605
14.4257 | -106.2631
14.4296 | -106.2715
14.4334 | -106.2787
14.4372 | -106.2856
14.4411 | -106.292
14.4449 | -106.2937
14.4488 | -106.296
14.4526 | -106.2982
14.4564 | -106.2993
14.4603 | -106.3003
14.4641 | -106.3004
14.468 | -106.2992
14.4718 | -106.2939
14.4756 | -106.2884
14.4795 | -106.2845
14.4833 | -106.2789
14.4872 | -106.2715
14.491 | -106.2621
14.4948 | -106.2513
14.4987 | -106.2373
14.5025 | -106.2225
14.5064 | -106.208
14.5102 | -106.1906
14.514 | -106.1722
14.5179 | -106.1562
14.5217 | -106.1397
14.5256 | -106.1251
14.5294 | -106.1111
14.5332 | -106.0957
14.5371 | -106.0806
14.5409 | -106.063
14.5448 | -106.0491
14.5486 | -106.0413
14.5524 | -106.0319
14.5563 | -106.0238
14.5601 | -106.018
14.564 | -106.0165
14.5678 | -106.0161
14.5716 | -106.0147
14.5755 | -106.0164
14.5793 | -106.0192
14.5832 | -106.0234
14.587 | -106.0291
14.5908 | -106.0363
14.5947 | -106.0444
14.5985 | -106.053
14.6024 | -106.0592
14.6062 | -106.0661
14.61 | -106.0754
14.6139 | -106.0848
14.6177 | -106.0959
14.6215 | -106.1056
14.6254 | -106.1162
14.6292 | -106.1288
14.6331 | -106.1416
14.6369 | -106.1538
14.6407 | -106.166
14.6446 | -106.1772
14.6484 | -106.1865
14.6523 | -106.1949
14.6561 | -106.2059
14.6599 | -106.2175
14.6638 | -106.2298
14.6676 | -106.2388
14.6715 | -106.2475
14.6753 | -106.2567
14.6791 | -106.261
14.683 | -106.2654
14.6868 | -106.2702
14.6907 | -106.2729
14.6945 | -106.2737
14.6983 | -106.2775
14.7022 | -106.2818
14.706 | -106.287
14.7099 | -106.2944
14.7137 | -106.3007
14.7175 | -106.3074
14.7214 | -106.318
14.7252 | -106.3297
14.7291 | -106.3378
14.7329 | -106.345
14.7367 | -106.3556
14.7406 | -106.3685
14.7444 | -106.3831
14.7483 | -106.4014
14.7521 | -106.4224
14.7559 | -106.4366
14.7598 | -106.4506
14.7636 | -106.4666
14.7675 | -106.4843
14.7713 | -106.4987
14.7751 | -106.5105
14.779 | -106.5228
14.7828 | -106.5286
14.7867 | -106.5346
14.7905 | -106.5391
14.7943 | -106.5369
14.7982 | -106.5389
14.802 | -106.5411
14.8059 | -106.5404
14.8097 | -106.5369
14.8135 | -106.5325
14.8174 | -106.5276
14.8212 | -106.5176
14.8251 | -106.5063
14.8289 | -106.4952
14.8327 | -106.4813
14.8366 | -106.4682
14.8404 | -106.4502
14.8443 | -106.4354
14.8481 | -106.4236
14.8519 | -106.4092
14.8558 | -106.396
14.8596 | -106.3844
14.8635 | -106.3727
14.8673 | -106.3622
14.8711 | -106.3528
14.875 | -106.3413
14.8788 | -106.331
14.8826 | -106.322
14.8865 | -106.3134
14.8903 | -106.3083
14.8942 | -106.3033
14.898 | -106.299
14.9018 | -106.2959
14.9057 | -106.2947
14.9095 | -106.293
14.9134 | -106.2926
14.9172 | -106.2923
14.921 | -106.2935
14.9249 | -106.2969
14.9287 | -106.2985
14.9326 | -106.2964
14.9364 | -106.2945
14.9402 | -106.2965
14.9441 | -106.3029
14.9479 | -106.3064
14.9518 | -106.3083
14.9556 | -106.3121
14.9594 | -106.3136
14.9633 | -106.3126
14.9671 | -106.3137
14.971 | -106.3156
14.9748 | -106.3122
14.9786 | -106.303
14.9825 | -106.2944
14.9863 | -106.2895
14.9902 | -106.2837
14.994 | -106.2785
14.9978 | -106.2763
15.0017 | -106.2717
15.0055 | -106.2674
15.0094 | -106.2636
15.0132 | -106.2517
15.017 | -106.241
15.0209 | -106.2291
15.0247 | -106.215
15.0286 | -106.201
15.0324 | -106.1891
15.0362 | -106.176
15.0401 | -106.1643
15.0439 | -106.1518
15.0478 | -106.1415
15.0516 | -106.1303
15.0554 | -106.1209
15.0593 | -106.1099
15.0631 | -106.0974
15.067 | -106.0886
15.0708 | -106.0812
15.0746 | -106.0775
15.0785 | -106.0706
15.0823 | -106.0652
15.0862 | -106.0632
15.09 | -106.0615
15.0938 | -106.0644
15.0977 | -106.0684
15.1015 | -106.0721
15.1054 | -106.077
15.1092 | -106.0828
15.113 | -106.0863
15.1169 | -106.0918
15.1207 | -106.0999
15.1245 | -106.1084
15.1284 | -106.1159
15.1322 | -106.1228
15.1361 | -106.1313
15.1399 | -106.14
15.1437 | -106.1491
15.1476 | -106.1606
15.1514 | -106.1721
15.1553 | -106.1799
15.1591 | -106.186
15.1629 | -106.1921
15.1668 | -106.1976
15.1706 | -106.2039
15.1745 | -106.2091
15.1783 | -106.2172
15.1821 | -106.2265
15.186 | -106.2315
15.1898 | -106.2392
15.1937 | -106.2494
15.1975 | -106.2619
15.2013 | -106.2738
15.2052 | -106.2836
15.209 | -106.295
15.2129 | -106.3031
15.2167 | -106.311
15.2205 | -106.3203
15.2244 | -106.3289
15.2282 | -106.3435
15.2321 | -106.361
15.2359 | -106.3784
15.2397 | -106.3968
15.2436 | -106.4147
15.2474 | -106.4325
15.2513 | -106.4466
15.2551 | -106.4613
15.2589 | -106.4747
15.2628 | -106.4907
15.2666 | -106.5065
15.2705 | -106.5212
15.2743 | -106.5328
15.2781 | -106.5416
15.282 | -106.551
15.2858 | -106.558
15.2897 | -106.563
15.2935 | -106.5619
15.2973 | -106.5629
15.3012 | -106.5643
15.305 | -106.5604
15.3089 | -106.5559
15.3127 | -106.5493
15.3165 | -106.5415
15.3204 | -106.5352
15.3242 | -106.5271
15.3281 | -106.5197
15.3319 | -106.513
15.3357 | -106.5036
15.3396 | -106.4929
15.3434 | -106.4836
15.3473 | -106.4775
15.3511 | -106.4715
15.3549 | -106.4639
15.3588 | -106.458
15.3626 | -106.4564
15.3665 | -106.4526
15.3703 | -106.4504
15.3741 | -106.4489
15.378 | -106.445
15.3818 | -106.4447
15.3856 | -106.4455
15.3895 | -106.4465
15.3933 | -106.4509
15.3972 | -106.4537
15.401 | -106.4565
15.4048 | -106.4608
15.4087 | -106.4655
15.4125 | -106.4706
15.4164 | -106.4781
15.4202 | -106.4834
15.424 | -106.4869
15.4279 | -106.4913
15.4317 | -106.4991
15.4356 | -106.5062
15.4394 | -106.511
15.4432 | -106.5148
15.4471 | -106.513
15.4509 | -106.5107
15.4548 | -106.5076
15.4586 | -106.5039
15.4624 | -106.4993
15.4663 | -106.4951
15.4701 | -106.4882
15.474 | -106.4802
15.4778 | -106.4711
15.4816 | -106.4579
15.4855 | -106.444
15.4893 | -106.4292
15.4932 | -106.4162
15.497 | -106.4011
15.5008 | -106.3834
15.5047 | -106.3674
15.5085 | -106.3531
15.5124 | -106.3402
15.5162 | -106.3297
15.52 | -106.3163
15.5239 | -106.3013
15.5277 | -106.292
15.5316 | -106.2856
15.5354 | -106.2786
15.5392 | -106.272
15.5431 | -106.2657
15.5469 | -106.262
15.5508 | -106.2599
15.5546 | -106.2589
15.5584 | -106.2593
15.5623 | -106.2592
15.5661 | -106.2608
15.57 | -106.2628
15.5738 | -106.2675
15.5776 | -106.2723
15.5815 | -106.2753
15.5853 | -106.2786
15.5892 | -106.2862
15.593 | -106.2919
15.5968 | -106.2994
15.6007 | -106.3102
15.6045 | -106.319
15.6084 | -106.3259
15.6122 | -106.3325
15.616 | -106.3391
15.6199 | -106.3451
15.6237 | -106.3528
15.6275 | -106.3605
15.6314 | -106.3636
15.6352 | -106.3642
15.6391 | -106.3643
15.6429 | -106.3675
15.6467 | -106.3691
15.6506 | -106.3706
15.6544 | -106.369
15.6583 | -106.3675
15.6621 | -106.3696
15.6659 | -106.3692
15.6698 | -106.369
15.6736 | -106.3659
15.6775 | -106.3624
15.6813 | -106.3599
15.6851 | -106.3587
15.689 | -106.355
15.6928 | -106.351
15.6967 | -106.3468
15.7005 | -106.3426
15.7043 | -106.3375
15.7082 | -106.3305
15.712 | -106.3263
15.7159 | -106.3238
15.7197 | -106.3218
15.7235 | -106.3173
15.7274 | -106.3128
15.7312 | -106.3083
15.7351 | -106.3059
15.7389 | -106.3041
15.7427 | -106.3051
15.7466 | -106.3098
15.7504 | -106.3119
15.7543 | -106.3111
15.7581 | -106.3113
15.7619 | -106.3132
15.7658 | -106.317
15.7696 | -106.3212
15.7735 | -106.3251
15.7773 | -106.3298
15.7811 | -106.335
15.785 | -106.3379
15.7888 | -106.3414
15.7927 | -106.3445
15.7965 | -106.3482
15.8003 | -106.3484
15.8042 | -106.3485
15.808 | -106.351
15.8119 | -106.3492
15.8157 | -106.3428
15.8195 | -106.3363
15.8234 | -106.3271
15.8272 | -106.3166
15.8311 | -106.3029
15.8349 | -106.2918
15.8387 | -106.2824
15.8426 | -106.2687
15.8464 | -106.2555
15.8503 | -106.239
15.8541 | -106.2241
15.8579 | -106.2132
15.8618 | -106.2021
15.8656 | -106.1891
15.8695 | -106.1726
15.8733 | -106.1626
15.8771 | -106.1512
15.881 | -106.1403
15.8848 | -106.1324
15.8886 | -106.1227
15.8925 | -106.1181
15.8963 | -106.1162
15.9002 | -106.1145
15.904 | -106.113
15.9078 | -106.1138
15.9117 | -106.1174
15.9155 | -106.1251
15.9194 | -106.136
15.9232 | -106.1491
15.927 | -106.1637
15.9309 | -106.1812
15.9347 | -106.2027
15.9386 | -106.2226
15.9424 | -106.2408
15.9462 | -106.2624
15.9501 | -106.2858
15.9539 | -106.3096
15.9578 | -106.3358
15.9616 | -106.3631
15.9654 | -106.3879
15.9693 | -106.412
15.9731 | -106.4371
15.977 | -106.4649
15.9808 | -106.4921
15.9846 | -106.5174
15.9885 | -106.5401
15.9923 | -106.5602
15.9962 | -106.5777
16 | -106.5924
16.0038 | -106.6074
16.0077 | -106.6228
16.0115 | -106.6386
16.0154 | -106.6495
16.0192 | -106.6582
16.023 | -106.6661
16.0269 | -106.676
16.0307 | -106.6839
16.0346 | -106.6895
16.0384 | -106.6972
16.0422 | -106.702
16.0461 | -106.706
16.0499 | -106.7112
16.0538 | -106.7154
16.0576 | -106.7178
16.0614 | -106.7121
16.0653 | -106.7111
16.0691 | -106.7083
16.073 | -106.7073
16.0768 | -106.706
16.0806 | -106.7053
16.0845 | -106.7016
16.0883 | -106.6937
16.0922 | -106.6871
16.096 | -106.6796
16.0998 | -106.6731
16.1037 | -106.6664
16.1075 | -106.6613
16.1114 | -106.6568
16.1152 | -106.646
16.119 | -106.637
16.1229 | -106.6301
16.1267 | -106.6232
16.1305 | -106.6188
16.1344 | -106.6139
16.1382 | -106.6052
16.1421 | -106.5982
16.1459 | -106.5901
16.1497 | -106.583
16.1536 | -106.5763
16.1574 | -106.5692
16.1613 | -106.5644
16.1651 | -106.5602
16.1689 | -106.5563
16.1728 | -106.5522
16.1766 | -106.5431
16.1805 | -106.5342
16.1843 | -106.5232
16.1881 | -106.5093
16.192 | -106.4944
16.1958 | -106.4823
16.1997 | -106.4698
16.2035 | -106.4589
16.2073 | -106.4473
16.2112 | -106.4344
16.215 | -106.4209
16.2189 | -106.4061
16.2227 | -106.3916
16.2265 | -106.3782
16.2304 | -106.3651
16.2342 | -106.352
16.2381 | -106.3364
16.2419 | -106.3239
16.2457 | -106.3151
16.2496 | -106.3054
16.2534 | -106.2938
16.2573 | -106.2793
16.2611 | -106.2678
16.2649 | -106.2567
16.2688 | -106.2462
16.2726 | -106.2379
16.2765 | -106.2292
16.2803 | -106.2185
16.2841 | -106.208
16.288 | -106.2028
16.2918 | -106.1983
16.2957 | -106.1939
16.2995 | -106.1897
16.3033 | -106.1819
16.3072 | -106.176
16.311 | -106.1723
16.3149 | -106.1699
16.3187 | -106.1685
16.3225 | -106.1666
16.3264 | -106.163
16.3302 | -106.1629
16.3341 | -106.1615
16.3379 | -106.1559
16.3417 | -106.152
16.3456 | -106.1501
16.3494 | -106.1506
16.3533 | -106.1502
16.3571 | -106.1473
16.3609 | -106.1459
16.3648 | -106.1451
16.3686 | -106.146
16.3725 | -106.1473
16.3763 | -106.1459
16.3801 | -106.1451
16.384 | -106.1461
16.3878 | -106.1447
16.3916 | -106.1446
16.3955 | -106.1442
16.3993 | -106.1447
16.4032 | -106.1474
16.407 | -106.1508
16.4108 | -106.1562
16.4147 | -106.1621
16.4185 | -106.1683
16.4224 | -106.1728
16.4262 | -106.1769
16.43 | -106.1798
16.4339 | -106.1849
16.4377 | -106.1912
16.4416 | -106.1989
16.4454 | -106.2093
16.4492 | -106.2226
16.4531 | -106.2364
16.4569 | -106.2495
16.4608 | -106.2632
16.4646 | -106.2778
16.4684 | -106.2898
16.4723 | -106.3045
16.4761 | -106.3196
16.48 | -106.3332
16.4838 | -106.3473
16.4876 | -106.3626
16.4915 | -106.382
16.4953 | -106.4004
16.4992 | -106.4178
16.503 | -106.4342
16.5068 | -106.4532
16.5107 | -106.4704
16.5145 | -106.4871
16.5184 | -106.5024
16.5222 | -106.5138
16.526 | -106.5228
16.5299 | -106.5327
16.5337 | -106.5435
16.5376 | -106.5553
16.5414 | -106.5643
16.5452 | -106.571
16.5491 | -106.5754
16.5529 | -106.5798
16.5568 | -106.5856
16.5606 | -106.5907
16.5644 | -106.5926
16.5683 | -106.5944
16.5721 | -106.5938
16.576 | -106.5909
16.5798 | -106.5886
16.5836 | -106.5862
16.5875 | -106.5826
16.5913 | -106.5794
16.5952 | -106.575
16.599 | -106.5703
16.6028 | -106.5667
16.6067 | -106.5652
16.6105 | -106.5616
16.6144 | -106.5537
16.6182 | -106.5464
16.622 | -106.5404
16.6259 | -106.5377
16.6297 | -106.534
16.6335 | -106.5283
16.6374 | -106.5234
16.6412 | -106.519
16.6451 | -106.5148
16.6489 | -106.5125
16.6527 | -106.5091
16.6566 | -106.5076
16.6604 | -106.5064
16.6643 | -106.505
16.6681 | -106.5052
16.6719 | -106.5038
16.6758 | -106.4986
16.6796 | -106.4951
16.6835 | -106.4951
16.6873 | -106.4944
16.6911 | -106.4953
16.695 | -106.4985
16.6988 | -106.5005
16.7027 | -106.5017
16.7065 | -106.5045
16.7103 | -106.5091
16.7142 | -106.5137
16.718 | -106.5209
16.7219 | -106.5292
16.7257 | -106.5394
16.7295 | -106.5476
16.7334 | -106.5574
16.7372 | -106.5682
16.7411 | -106.5779
16.7449 | -106.5862
16.7487 | -106.5981
16.7526 | -106.6104
16.7564 | -106.6231
16.7603 | -106.6362
16.7641 | -106.6517
16.7679 | -106.6639
16.7718 | -106.674
16.7756 | -106.6839
16.7795 | -106.6952
16.7833 | -106.7066
16.7871 | -106.7167
16.791 | -106.7246
16.7948 | -106.7313
16.7987 | -106.7382
16.8025 | -106.7422
16.8063 | -106.7466
16.8102 | -106.7524
16.814 | -106.757
16.8179 | -106.7569
16.8217 | -106.7561
16.8255 | -106.7543
16.8294 | -106.7523
16.8332 | -106.7491
16.8371 | -106.7453
16.8409 | -106.7403
16.8447 | -106.7348
16.8486 | -106.7326
16.8524 | -106.7333
16.8563 | -106.7319
16.8601 | -106.7278
16.8639 | -106.7227
16.8678 | -106.7163
16.8716 | -106.7122
16.8755 | -106.711
16.8793 | -106.7081
16.8831 | -106.7043
16.887 | -106.704
16.8908 | -106.7043
16.8946 | -106.7041
16.8985 | -106.7024
16.9023 | -106.7046
16.9062 | -106.7081
16.91 | -106.7106
16.9138 | -106.7132
16.9177 | -106.7185
16.9215 | -106.724
16.9254 | -106.7289
16.9292 | -106.735
16.933 | -106.74
16.9369 | -106.7416
16.9407 | -106.7464
16.9446 | -106.7521
16.9484 | -106.7575
16.9522 | -106.762
16.9561 | -106.7631
16.9599 | -106.7614
16.9638 | -106.7587
16.9676 | -106.7582
16.9714 | -106.7582
16.9753 | -106.7551
16.9791 | -106.7534
16.983 | -106.7535
16.9868 | -106.7517
16.9906 | -106.7486
16.9945 | -106.7465
16.9983 | -106.7458
17.0022 | -106.7456
17.006 | -106.7434
17.0098 | -106.7418
17.0137 | -106.7411
17.0175 | -106.7419
17.0214 | -106.7421
17.0252 | -106.7423
17.029 | -106.7445
17.0329 | -106.7473
17.0367 | -106.7531
17.0406 | -106.7577
17.0444 | -106.7601
17.0482 | -106.7587
17.0521 | -106.7596
17.0559 | -106.761
17.0598 | -106.7636
17.0636 | -106.7628
17.0674 | -106.7617
17.0713 | -106.7607
17.0751 | -106.7606
17.079 | -106.7604
17.0828 | -106.7602
17.0866 | -106.7578
17.0905 | -106.7546
17.0943 | -106.7496
17.0982 | -106.7444
17.102 | -106.7391
17.1058 | -106.7339
17.1097 | -106.7274
17.1135 | -106.7171
17.1174 | -106.7057
17.1212 | -106.6962
17.125 | -106.6863
17.1289 | -106.6753
17.1327 | -106.6649
17.1365 | -106.6578
17.1404 | -106.6505
17.1442 | -106.6405
17.1481 | -106.6288
17.1519 | -106.6153
17.1557 | -106.606
17.1596 | -106.5985
17.1634 | -106.5905
17.1673 | -106.5799
17.1711 | -106.5756
17.1749 | -106.5677
17.1788 | -106.5625
17.1826 | -106.5611
17.1865 | -106.5581
17.1903 | -106.5535
17.1941 | -106.548
17.198 | -106.5452
17.2018 | -106.5467
17.2057 | -106.5492
17.2095 | -106.5508
17.2133 | -106.55
17.2172 | -106.5501
17.221 | -106.5498
17.2249 | -106.5507
17.2287 | -106.5522
17.2325 | -106.5538
17.2364 | -106.5555
17.2402 | -106.5569
17.2441 | -106.5541
17.2479 | -106.5535
17.2517 | -106.5579
17.2556 | -106.5614
17.2594 | -106.5649
17.2633 | -106.5679
17.2671 | -106.5698
17.2709 | -106.5731
17.2748 | -106.5768
17.2786 | -106.5822
17.2825 | -106.5835
17.2863 | -106.5864
17.2901 | -106.5914
17.294 | -106.599
17.2978 | -106.6052
17.3017 | -106.6085
17.3055 | -106.6103
17.3093 | -106.614
17.3132 | -106.6168
17.317 | -106.6188
17.3209 | -106.6196
17.3247 | -106.6213
17.3285 | -106.6227
17.3324 | -106.6266
17.3362 | -106.629
17.3401 | -106.6326
17.3439 | -106.6368
17.3477 | -106.6398
17.3516 | -106.6424
17.3554 | -106.6469
17.3593 | -106.6502
17.3631 | -106.6522
17.3669 | -106.6537
17.3708 | -106.6575
17.3746 | -106.6598
17.3785 | -106.6608
17.3823 | -106.66
17.3861 | -106.6612
17.39 | -106.6636
17.3938 | -106.6687
17.3976 | -106.6756
17.4015 | -106.679
17.4053 | -106.6812
17.4092 | -106.6853
17.413 | -106.6919
17.4168 | -106.6955
17.4207 | -106.6992
17.4245 | -106.7024
17.4284 | -106.7069
17.4322 | -106.7112
17.436 | -106.7151
17.4399 | -106.7213
17.4437 | -106.7308
17.4476 | -106.7374
17.4514 | -106.7419
17.4552 | -106.7503
17.4591 | -106.7607
17.4629 | -106.7709
17.4668 | -106.7779
17.4706 | -106.7884
17.4744 | -106.7953
17.4783 | -106.7999
17.4821 | -106.8057
17.486 | -106.8107
17.4898 | -106.8147
17.4936 | -106.8192
17.4975 | -106.8219
17.5013 | -106.8191
17.5052 | -106.8167
17.509 | -106.8128
17.5128 | -106.8073
17.5167 | -106.8005
17.5205 | -106.7895
17.5244 | -106.7803
17.5282 | -106.7685
17.532 | -106.7563
17.5359 | -106.7461
17.5397 | -106.7352
17.5436 | -106.7237
17.5474 | -106.7102
17.5512 | -106.6962
17.5551 | -106.681
17.5589 | -106.67
17.5628 | -106.6556
17.5666 | -106.6406
17.5704 | -106.6283
17.5743 | -106.6155
17.5781 | -106.6013
17.582 | -106.5901
17.5858 | -106.5796
17.5896 | -106.5691
17.5935 | -106.5589
17.5973 | -106.5494
17.6012 | -106.5431
17.605 | -106.5391
17.6088 | -106.5384
17.6127 | -106.5385
17.6165 | -106.5387
17.6203 | -106.5388
17.6242 | -106.5417
17.628 | -106.545
17.6319 | -106.5508
17.6357 | -106.5577
17.6395 | -106.5671
17.6434 | -106.5746
17.6472 | -106.5821
17.6511 | -106.5912
17.6549 | -106.604
17.6587 | -106.6173
17.6626 | -106.6305
17.6664 | -106.6481
17.6703 | -106.6648
17.6741 | -106.6797
17.6779 | -106.6939
17.6818 | -106.7082
17.6856 | -106.7234
17.6895 | -106.7384
17.6933 | -106.7516
17.6971 | -106.7656
17.701 | -106.7802
17.7048 | -106.7932
17.7087 | -106.8052
17.7125 | -106.8166
17.7163 | -106.8326
17.7202 | -106.8453
17.724 | -106.8551
17.7279 | -106.8679
17.7317 | -106.8813
17.7355 | -106.8947
17.7394 | -106.906
17.7432 | -106.9155
17.7471 | -106.9244
17.7509 | -106.9307
17.7547 | -106.9396
17.7586 | -106.9507
17.7624 | -106.9609
17.7663 | -106.9701
17.7701 | -106.9782
17.7739 | -106.9883
17.7778 | -106.9943
17.7816 | -106.999
17.7855 | -107.0054
17.7893 | -107.009
17.7931 | -107.0119
17.797 | -107.0124
17.8008 | -107.0121
17.8047 | -107.0136
17.8085 | -107.0162
17.8123 | -107.014
17.8162 | -107.0138
17.82 | -107.0106
17.8239 | -107.0044
17.8277 | -106.9975
17.8315 | -106.9903
17.8354 | -106.9813
17.8392 | -106.9725
17.8431 | -106.9626
17.8469 | -106.9532
17.8507 | -106.9424
17.8546 | -106.9281
17.8584 | -106.9121
17.8623 | -106.8971
17.8661 | -106.8831
17.8699 | -106.8691
17.8738 | -106.8523
17.8776 | -106.835
17.8814 | -106.821
17.8853 | -106.8057
17.8891 | -106.79
17.893 | -106.7746
17.8968 | -106.763
17.9006 | -106.7526
17.9045 | -106.7433
17.9083 | -106.7336
17.9122 | -106.7257
17.916 | -106.7195
17.9198 | -106.7114
17.9237 | -106.705
17.9275 | -106.7012
17.9314 | -106.7
17.9352 | -106.6998
17.939 | -106.7001
17.9429 | -106.7015
17.9467 | -106.7025
17.9506 | -106.7037
17.9544 | -106.7067
17.9582 | -106.7108
17.9621 | -106.7152
17.9659 | -106.7188
17.9698 | -106.7232
17.9736 | -106.7272
17.9774 | -106.7314
17.9813 | -106.7366
17.9851 | -106.74
17.989 | -106.7444
17.9928 | -106.7474
17.9966 | -106.7522
18.0005 | -106.7579
18.0043 | -106.7634
18.0082 | -106.7691
18.012 | -106.7737
18.0158 | -106.775
18.0197 | -106.7776
18.0235 | -106.7814
18.0274 | -106.788
18.0312 | -106.7919
18.035 | -106.7958
18.0389 | -106.802
18.0427 | -106.8062
18.0466 | -106.8067
18.0504 | -106.8064
18.0542 | -106.8084
18.0581 | -106.8125
18.0619 | -106.8155
18.0658 | -106.8178
18.0696 | -106.8184
18.0734 | -106.8161
18.0773 | -106.8152
18.0811 | -106.8149
18.085 | -106.8144
18.0888 | -106.813
18.0926 | -106.8084
18.0965 | -106.8043
18.1003 | -106.8016
18.1042 | -106.7985
18.108 | -106.7937
18.1118 | -106.7892
18.1157 | -106.7846
18.1195 | -106.7779
18.1233 | -106.7723
18.1272 | -106.7668
18.131 | -106.7646
18.1349 | -106.7602
18.1387 | -106.7538
18.1425 | -106.751
18.1464 | -106.7487
18.1502 | -106.7443
18.1541 | -106.7389
18.1579 | -106.7339
18.1617 | -106.7305
18.1656 | -106.7275
18.1694 | -106.724
18.1733 | -106.7206
18.1771 | -106.7186
18.1809 | -106.7175
18.1848 | -106.7177
18.1886 | -106.7151
18.1925 | -106.7159
18.1963 | -106.7178
18.2001 | -106.718
18.204 | -106.7183
18.2078 | -106.7199
18.2117 | -106.722
18.2155 | -106.7236
18.2193 | -106.7268
18.2232 | -106.73
18.227 | -106.7321
18.2309 | -106.7365
18.2347 | -106.7408
18.2385 | -106.7438
18.2424 | -106.747
18.2462 | -106.7495
18.2501 | -106.7513
18.2539 | -106.7517
18.2577 | -106.7523
18.2616 | -106.7553
18.2654 | -106.7569
18.2693 | -106.7568
18.2731 | -106.7584
18.2769 | -106.7564
18.2808 | -106.7561
18.2846 | -106.7538
18.2885 | -106.7511
18.2923 | -106.7485
18.2961 | -106.747
18.3 | -106.7464
18.3038 | -106.7448
18.3077 | -106.7455
18.3115 | -106.7476
18.3153 | -106.7466
18.3192 | -106.7443
18.323 | -106.7431
18.3269 | -106.7426
18.3307 | -106.7442
18.3345 | -106.749
18.3384 | -106.7511
18.3422 | -106.753
18.3461 | -106.7556
18.3499 | -106.7611
18.3537 | -106.7693
18.3576 | -106.7771
18.3614 | -106.7842
18.3653 | -106.7903
18.3691 | -106.7991
18.3729 | -106.8085
18.3768 | -106.8169
18.3806 | -106.8266
18.3844 | -106.8371
18.3883 | -106.8485
18.3921 | -106.8591
18.396 | -106.8724
18.3998 | -106.8858
18.4036 | -106.8953
18.4075 | -106.9034
18.4113 | -106.913
18.4152 | -106.9246
18.419 | -106.9376
18.4228 | -106.9497
18.4267 | -106.9612
18.4305 | -106.971
18.4344 | -106.9805
18.4382 | -106.9939
18.442 | -107.0062
18.4459 | -107.0152
18.4497 | -107.0255
18.4536 | -107.0339
18.4574 | -107.0436
18.4612 | -107.0544
18.4651 | -107.0671
18.4689 | -107.0787
18.4728 | -107.0889
18.4766 | -107.0976
18.4804 | -107.1082
18.4843 | -107.1199
18.4881 | -107.1304
18.492 | -107.1407
18.4958 | -107.151
18.4996 | -107.1615
18.5035 | -107.1706
18.5073 | -107.1772
18.5112 | -107.1812
18.515 | -107.1861
18.5188 | -107.189
18.5227 | -107.1895
18.5265 | -107.1898
18.5304 | -107.1898
18.5342 | -107.1884
18.538 | -107.188
18.5419 | -107.1847
18.5457 | -107.1805
18.5496 | -107.1745
18.5534 | -107.1674
18.5572 | -107.1616
18.5611 | -107.1536
18.5649 | -107.1415
18.5688 | -107.1314
18.5726 | -107.1224
18.5764 | -107.111
18.5803 | -107.1012
18.5841 | -107.0909
18.588 | -107.0815
18.5918 | -107.0706
18.5956 | -107.0595
18.5995 | -107.049
18.6033 | -107.0427
18.6072 | -107.0355
18.611 | -107.0289
18.6148 | -107.0233
18.6187 | -107.0168
18.6225 | -107.0114
18.6263 | -107.0106
18.6302 | -107.0113
18.634 | -107.0114
18.6379 | -107.0138
18.6417 | -107.0186
18.6455 | -107.024
18.6494 | -107.0268
18.6532 | -107.0301
18.6571 | -107.0357
18.6609 | -107.0445
18.6647 | -107.0524
18.6686 | -107.061
18.6724 | -107.0711
18.6763 | -107.0783
18.6801 | -107.0793
18.6839 | -107.084
18.6878 | -107.0901
18.6916 | -107.0951
18.6955 | -107.1003
18.6993 | -107.1067
18.7031 | -107.1087
18.707 | -107.1101
18.7108 | -107.1141
18.7147 | -107.1172
18.7185 | -107.1179
18.7223 | -107.1152
18.7262 | -107.111
18.73 | -107.1055
18.7339 | -107.0981
18.7377 | -107.0906
18.7415 | -107.0842
18.7454 | -107.0734
18.7492 | -107.0607
18.7531 | -107.0472
18.7569 | -107.0301
18.7607 | -107.0133
18.7646 | -106.9982
18.7684 | -106.9813
18.7723 | -106.965
18.7761 | -106.9493
18.7799 | -106.9316
18.7838 | -106.9141
18.7876 | -106.896
18.7915 | -106.8773
18.7953 | -106.8558
18.7991 | -106.8347
18.803 | -106.8147
18.8068 | -106.7952
18.8107 | -106.7759
18.8145 | -106.7601
18.8183 | -106.7436
18.8222 | -106.7263
18.826 | -106.7108
18.8299 | -106.6978
18.8337 | -106.6851
18.8375 | -106.6698
18.8414 | -106.6539
18.8452 | -106.6403
18.8491 | -106.6292
18.8529 | -106.6184
18.8567 | -106.6072
18.8606 | -106.5966
18.8644 | -106.5876
18.8683 | -106.58
18.8721 | -106.5717
18.8759 | -106.5658
18.8798 | -106.5609
18.8836 | -106.5571
18.8874 | -106.5535
18.8913 | -106.5518
18.8951 | -106.5493
18.899 | -106.5494
18.9028 | -106.552
18.9066 | -106.5538
18.9105 | -106.5554
18.9143 | -106.5579
18.9182 | -106.5653
18.922 | -106.5728
18.9258 | -106.5803
18.9297 | -106.5861
18.9335 | -106.5931
18.9374 | -106.6018
18.9412 | -106.6128
18.945 | -106.6261
18.9489 | -106.6418
18.9527 | -106.6522
18.9566 | -106.6621
18.9604 | -106.6749
18.9642 | -106.6896
18.9681 | -106.7046
18.9719 | -106.718
18.9758 | -106.7312
18.9796 | -106.7438
18.9834 | -106.7558
18.9873 | -106.7686
18.9911 | -106.7812
18.995 | -106.7927
18.9988 | -106.804
19.0026 | -106.8182
19.0065 | -106.8309
19.0103 | -106.8425
19.0142 | -106.8543
19.018 | -106.8671
19.0218 | -106.879
19.0257 | -106.8914
19.0295 | -106.9017
19.0334 | -106.9093
19.0372 | -106.9174
19.041 | -106.9238
19.0449 | -106.9278
19.0487 | -106.9317
19.0526 | -106.9356
19.0564 | -106.9377
19.0602 | -106.9407
19.0641 | -106.9443
19.0679 | -106.9454
19.0718 | -106.9466
19.0756 | -106.9473
19.0794 | -106.947
19.0833 | -106.9437
19.0871 | -106.9425
19.091 | -106.9401
19.0948 | -106.9371
19.0986 | -106.9333
19.1025 | -106.9331
19.1063 | -106.931
19.1102 | -106.9293
19.114 | -106.9257
19.1178 | -106.9224
19.1217 | -106.919
19.1255 | -106.9151
19.1293 | -106.9114
19.1332 | -106.9088
19.137 | -106.9045
19.1409 | -106.9011
19.1447 | -106.8996
19.1485 | -106.8964
19.1524 | -106.8942
19.1562 | -106.8915
19.1601 | -106.8912
19.1639 | -106.89
19.1677 | -106.8855
19.1716 | -106.8806
19.1754 | -106.8786
19.1793 | -106.8794
19.1831 | -106.8801
19.1869 | -106.8791
19.1908 | -106.8767
19.1946 | -106.8772
19.1985 | -106.8782
19.2023 | -106.8783
19.2061 | -106.8783
19.21 | -106.8784
19.2138 | -106.8749
19.2177 | -106.8723
19.2215 | -106.8708
19.2253 | -106.8717
19.2292 | -106.8717
19.233 | -106.872
19.2369 | -106.8716
19.2407 | -106.8708
19.2445 | -106.8745
19.2484 | -106.877
19.2522 | -106.8784
19.2561 | -106.88
19.2599 | -106.8818
19.2637 | -106.8838
19.2676 | -106.8883
19.2714 | -106.8904
19.2753 | -106.891
19.2791 | -106.8926
19.2829 | -106.8949
19.2868 | -106.8985
19.2906 | -106.9003
19.2945 | -106.9004
19.2983 | -106.8998
19.3021 | -106.901
19.306 | -106.9008
19.3098 | -106.9012
19.3137 | -106.9023
19.3175 | -106.9048
19.3213 | -106.9035
19.3252 | -106.9023
19.329 | -106.9019
19.3329 | -106.9018
19.3367 | -106.9016
19.3405 | -106.8995
19.3444 | -106.8939
19.3482 | -106.89
19.3521 | -106.8859
19.3559 | -106.8816
19.3597 | -106.876
19.3636 | -106.8718
19.3674 | -106.8667
19.3712 | -106.8633
19.3751 | -106.8585
19.3789 | -106.8556
19.3828 | -106.8531
19.3866 | -106.8465
19.3904 | -106.8436
19.3943 | -106.8387
19.3981 | -106.8312
19.402 | -106.8274
19.4058 | -106.8247
19.4096 | -106.8212
19.4135 | -106.8183
19.4173 | -106.8133
19.4212 | -106.8075
19.425 | -106.8055
19.4288 | -106.804
19.4327 | -106.8012
19.4365 | -106.7984
19.4404 | -106.798
19.4442 | -106.7989
19.448 | -106.8009
19.4519 | -106.8024
19.4557 | -106.8011
19.4596 | -106.8001
19.4634 | -106.8013
19.4672 | -106.8041
19.4711 | -106.8062
19.4749 | -106.8088
19.4788 | -106.813
19.4826 | -106.8166
19.4864 | -106.8214
19.4903 | -106.8274
19.4941 | -106.8291
19.498 | -106.8317
19.5018 | -106.8346
19.5056 | -106.8377
19.5095 | -106.8384
19.5133 | -106.8385
19.5172 | -106.8385
19.521 | -106.838
19.5248 | -106.8403
19.5287 | -106.843
19.5325 | -106.8451
19.5364 | -106.8476
19.5402 | -106.8511
19.544 | -106.8514
19.5479 | -106.8499
19.5517 | -106.8475
19.5556 | -106.8468
19.5594 | -106.8466
19.5632 | -106.846
19.5671 | -106.8472
19.5709 | -106.8483
19.5748 | -106.849
19.5786 | -106.8487
19.5824 | -106.8477
19.5863 | -106.8461
19.5901 | -106.8437
19.594 | -106.8427
19.5978 | -106.8416
19.6016 | -106.8367
19.6055 | -106.8337
19.6093 | -106.8307
19.6132 | -106.8289
19.617 | -106.8259
19.6208 | -106.8233
19.6247 | -106.8218
19.6285 | -106.82
19.6323 | -106.8168
19.6362 | -106.8124
19.64 | -106.8068
19.6439 | -106.802
19.6477 | -106.7976
19.6515 | -106.792
19.6554 | -106.7862
19.6592 | -106.7823
19.6631 | -106.782
19.6669 | -106.7804
19.6707 | -106.7765
19.6746 | -106.7727
19.6784 | -106.7672
19.6823 | -106.7602
19.6861 | -106.7554
19.6899 | -106.753
19.6938 | -106.7514
19.6976 | -106.7491
19.7015 | -106.7462
19.7053 | -106.7414
19.7091 | -106.7367
19.713 | -106.7338
19.7168 | -106.7319
19.7207 | -106.7304
19.7245 | -106.7316
19.7283 | -106.7315
19.7322 | -106.7295
19.736 | -106.7308
19.7399 | -106.7332
19.7437 | -106.734
19.7475 | -106.7326
19.7514 | -106.7322
19.7552 | -106.7323
19.7591 | -106.7336
19.7629 | -106.7358
19.7667 | -106.7351
19.7706 | -106.7355
19.7744 | -106.7401
19.7783 | -106.7445
19.7821 | -106.7475
19.7859 | -106.7491
19.7898 | -106.7523
19.7936 | -106.7571
19.7975 | -106.7618
19.8013 | -106.7646
19.8051 | -106.7691
19.809 | -106.7759
19.8128 | -106.784
19.8167 | -106.7915
19.8205 | -106.7985
19.8243 | -106.8088
19.8282 | -106.8197
19.832 | -106.8315
19.8359 | -106.8418
19.8397 | -106.8509
19.8435 | -106.8585
19.8474 | -106.867
19.8512 | -106.8783
19.8551 | -106.89
19.8589 | -106.8994
19.8627 | -106.9086
19.8666 | -106.9167
19.8704 | -106.9227
19.8742 | -106.9264
19.8781 | -106.9282
19.8819 | -106.9313
19.8858 | -106.9294
19.8896 | -106.9282
19.8934 | -106.9273
19.8973 | -106.9274
19.9011 | -106.9255
19.905 | -106.9192
19.9088 | -106.9127
19.9126 | -106.9073
19.9165 | -106.9003
19.9203 | -106.893
19.9242 | -106.8857
19.928 | -106.8764
19.9318 | -106.8673
19.9357 | -106.8585
19.9395 | -106.8474
19.9434 | -106.8373
19.9472 | -106.8246
19.951 | -106.8153
19.9549 | -106.8077
19.9587 | -106.8
19.9626 | -106.7959
19.9664 | -106.7908
19.9702 | -106.7821
19.9741 | -106.7725
19.9779 | -106.7634
19.9818 | -106.7577
19.9856 | -106.7531
19.9894 | -106.7482
19.9933 | -106.7443
19.9971 | -106.7398
20.001 | -106.7345
20.0048 | -106.7294
20.0086 | -106.7224
20.0125 | -106.717
20.0163 | -106.7144
20.0202 | -106.7083
20.024 | -106.7008
20.0278 | -106.697
20.0317 | -106.6942
20.0355 | -106.6893
20.0394 | -106.6849
20.0432 | -106.6798
20.047 | -106.6767
20.0509 | -106.6737
20.0547 | -106.6675
20.0586 | -106.6608
20.0624 | -106.6515
20.0662 | -106.6437
20.0701 | -106.6393
20.0739 | -106.6344
20.0778 | -106.63
20.0816 | -106.6252
20.0854 | -106.6203
20.0893 | -106.6146
20.0931 | -106.6071
20.097 | -106.6
20.1008 | -106.5943
20.1046 | -106.5898
20.1085 | -106.5838
20.1123 | -106.5802
20.1162 | -106.5766
20.12 | -106.5741
20.1238 | -106.5727
20.1277 | -106.5734
20.1315 | -106.5741
20.1353 | -106.5735
20.1392 | -106.5753
20.143 | -106.5809
20.1469 | -106.5861
20.1507 | -106.588
20.1545 | -106.592
20.1584 | -106.5975
20.1622 | -106.6048
20.1661 | -106.6123
20.1699 | -106.6206
20.1737 | -106.6262
20.1776 | -106.6311
20.1814 | -106.6366
20.1853 | -106.6421
20.1891 | -106.647
20.1929 | -106.6521
20.1968 | -106.6563
20.2006 | -106.6618
20.2045 | -106.6664
20.2083 | -106.6675
20.2121 | -106.6694
20.216 | -106.6722
20.2198 | -106.6713
20.2237 | -106.6704
20.2275 | -106.6651
20.2313 | -106.6613
20.2352 | -106.6601
20.239 | -106.6549
20.2429 | -106.6482
20.2467 | -106.6401
20.2505 | -106.6284
20.2544 | -106.6165
20.2582 | -106.6056
20.2621 | -106.594
20.2659 | -106.5788
20.2697 | -106.5611
20.2736 | -106.5446
20.2774 | -106.5276
20.2813 | -106.5089
20.2851 | -106.4925
20.2889 | -106.4764
20.2928 | -106.4577
20.2966 | -106.4417
20.3005 | -106.4247
20.3043 | -106.4061
20.3081 | -106.391
20.312 | -106.3784
20.3158 | -106.3622
20.3197 | -106.3471
20.3235 | -106.3321
20.3273 | -106.3162
20.3312 | -106.3008
20.335 | -106.2897
20.3389 | -106.2802
20.3427 | -106.2715
20.3465 | -106.2629
20.3504 | -106.2571
20.3542 | -106.2512
20.3581 | -106.2435
20.3619 | -106.2338
20.3657 | -106.2276
20.3696 | -106.2244
20.3734 | -106.222
20.3773 | -106.2188
20.3811 | -106.217
20.3849 | -106.2174
20.3888 | -106.2172
20.3926 | -106.2179
20.3964 | -106.2189
20.4003 | -106.2209
20.4041 | -106.2238
20.408 | -106.2249
20.4118 | -106.2255
20.4156 | -106.2272
20.4195 | -106.228
20.4233 | -106.2291
20.4272 | -106.229
20.431 | -106.2285
20.4348 | -106.2271
20.4387 | -106.2281
20.4425 | -106.2272
20.4464 | -106.2263
20.4502 | -106.2244
20.454 | -106.2196
20.4579 | -106.2156
20.4617 | -106.2123
20.4656 | -106.207
20.4694 | -106.2025
20.4732 | -106.1979
20.4771 | -106.1941
20.4809 | -106.1885
20.4848 | -106.18
20.4886 | -106.1716
20.4924 | -106.1659
20.4963 | -106.1598
20.5001 | -106.1555
20.504 | -106.1475
20.5078 | -106.1396
20.5116 | -106.1329
20.5155 | -106.1251
20.5193 | -106.1189
20.5232 | -106.1113
20.527 | -106.1028
20.5308 | -106.0959
20.5347 | -106.0889
20.5385 | -106.0832
20.5424 | -106.0798
20.5462 | -106.0746
20.55 | -106.0692
20.5539 | -106.0666
20.5577 | -106.0658
20.5616 | -106.064
20.5654 | -106.0609
20.5692 | -106.0586
20.5731 | -106.0573
20.5769 | -106.0539
20.5808 | -106.0508
20.5846 | -106.0502
20.5884 | -106.0493
20.5923 | -106.0497
20.5961 | -106.0511
20.6 | -106.0518
20.6038 | -106.0522
20.6076 | -106.0531
20.6115 | -106.0526
20.6153 | -106.055
20.6192 | -106.0582
20.623 | -106.0595
20.6268 | -106.0575
20.6307 | -106.0585
20.6345 | -106.0605
20.6383 | -106.0608
20.6422 | -106.0583
20.646 | -106.0579
20.6499 | -106.0575
20.6537 | -106.0576
20.6575 | -106.0566
20.6614 | -106.0548
20.6652 | -106.0539
20.6691 | -106.0527
20.6729 | -106.0526
20.6767 | -106.0501
20.6806 | -106.0468
20.6844 | -106.044
20.6883 | -106.0423
20.6921 | -106.037
20.6959 | -106.0333
20.6998 | -106.0317
20.7036 | -106.03
20.7075 | -106.0265
20.7113 | -106.023
20.7151 | -106.0206
20.719 | -106.0184
20.7228 | -106.0145
20.7267 | -106.0133
20.7305 | -106.0111
20.7343 | -106.0073
20.7382 | -106.0007
20.742 | -105.9946
20.7459 | -105.9896
20.7497 | -105.985
20.7535 | -105.9786
20.7574 | -105.9729
20.7612 | -105.9669
20.7651 | -105.9607
20.7689 | -105.9551
20.7727 | -105.9506
20.7766 | -105.9483
20.7804 | -105.9461
20.7843 | -105.9427
20.7881 | -105.9398
20.7919 | -105.935
20.7958 | -105.9319
20.7996 | -105.9317
20.8035 | -105.9307
20.8073 | -105.9309
20.8111 | -105.9306
20.815 | -105.9307
20.8188 | -105.9305
20.8227 | -105.9272
20.8265 | -105.9255
20.8303 | -105.9257
20.8342 | -105.9247
20.838 | -105.9223
20.8419 | -105.9225
20.8457 | -105.9238
20.8495 | -105.9226
20.8534 | -105.9216
20.8572 | -105.923
20.8611 | -105.9241
20.8649 | -105.9255
20.8687 | -105.9266
20.8726 | -105.9313
20.8764 | -105.9344
20.8803 | -105.9359
20.8841 | -105.9382
20.8879 | -105.9413
20.8918 | -105.9421
20.8956 | -105.9424
20.8994 | -105.9457
20.9033 | -105.9501
20.9071 | -105.953
20.911 | -105.9558
20.9148 | -105.9586
20.9186 | -105.9612
20.9225 | -105.9652
20.9263 | -105.9677
20.9302 | -105.9713
20.934 | -105.974
20.9378 | -105.9751
20.9417 | -105.9727
20.9455 | -105.9736
20.9494 | -105.9737
20.9532 | -105.971
20.957 | -105.965
20.9609 | -105.9574
20.9647 | -105.9498
20.9686 | -105.9406
20.9724 | -105.9322
20.9762 | -105.924
20.9801 | -105.9141
20.9839 | -105.8998
20.9878 | -105.8881
20.9916 | -105.8742
20.9954 | -105.8584
20.9993 | -105.842
21.0031 | -105.8271
21.007 | -105.8116
21.0108 | -105.7952
21.0146 | -105.7769
21.0185 | -105.7577
21.0223 | -105.7386
21.0262 | -105.7216
21.03 | -105.7056
21.0338 | -105.6883
21.0377 | -105.6706
21.0415 | -105.6546
21.0454 | -105.6383
21.0492 | -105.6289
21.053 | -105.6158
21.0569 | -105.6034
21.0607 | -105.5923
21.0646 | -105.5827
21.0684 | -105.573
21.0722 | -105.5642
21.0761 | -105.5574
21.0799 | -105.5508
21.0838 | -105.5461
21.0876 | -105.5431
21.0914 | -105.5435
21.0953 | -105.543
21.0991 | -105.5422
21.103 | -105.5438
21.1068 | -105.5489
21.1106 | -105.5552
21.1145 | -105.5612
21.1183 | -105.5677
21.1222 | -105.5776
21.126 | -105.5865
21.1298 | -105.5946
21.1337 | -105.6026
21.1375 | -105.6132
21.1413 | -105.6259
21.1452 | -105.638
21.149 | -105.6501
21.1529 | -105.6614
21.1567 | -105.6724
21.1605 | -105.6847
21.1644 | -105.697
21.1682 | -105.7105
21.1721 | -105.7229
21.1759 | -105.7372
21.1797 | -105.7531
21.1836 | -105.7667
21.1874 | -105.78
21.1913 | -105.7944
21.1951 | -105.8089
21.1989 | -105.8218
21.2028 | -105.8349
21.2066 | -105.8479
21.2105 | -105.8611
21.2143 | -105.8749
21.2181 | -105.8891
21.222 | -105.8993
21.2258 | -105.9083
21.2297 | -105.9186
21.2335 | -105.9304
21.2373 | -105.9443
21.2412 | -105.9561
21.245 | -105.9631
21.2489 | -105.9704
21.2527 | -105.9782
21.2565 | -105.9866
21.2604 | -105.9945
21.2642 | -106.0017
21.2681 | -106.0079
21.2719 | -106.0126
21.2757 | -106.0185
21.2796 | -106.0222
21.2834 | -106.0286
21.2873 | -106.0326
21.2911 | -106.0339
21.2949 | -106.0371
21.2988 | -106.0411
21.3026 | -106.0422
21.3065 | -106.0439
21.3103 | -106.044
21.3141 | -106.0445
21.318 | -106.0483
21.3218 | -106.0516
21.3257 | -106.0524
21.3295 | -106.0516
21.3333 | -106.0512
21.3372 | -106.052
21.341 | -106.0502
21.3449 | -106.046
21.3487 | -106.0427
21.3525 | -106.0408
21.3564 | -106.0375
21.3602 | -106.0356
21.3641 | -106.0336
21.3679 | -106.0299
21.3717 | -106.0246
21.3756 | -106.0207
21.3794 | -106.0179
21.3832 | -106.0144
21.3871 | -106.0098
21.3909 | -106.0077
21.3948 | -106.0056
21.3986 | -106.0028
21.4024 | -106.0027
21.4063 | -106.0018
21.4101 | -105.9985
21.414 | -105.996
21.4178 | -105.9955
21.4216 | -105.9951
21.4255 | -105.9973
21.4293 | -105.9994
21.4332 | -106.0003
21.437 | -106.003
21.4408 | -106.0052
21.4447 | -106.0055
21.4485 | -106.0061
21.4524 | -106.0075
21.4562 | -106.0086
21.46 | -106.0113
21.4639 | -106.0132
21.4677 | -106.0159
21.4716 | -106.0168
21.4754 | -106.0189
21.4792 | -106.0214
21.4831 | -106.0229
21.4869 | -106.0261
21.4908 | -106.0282
21.4946 | -106.0306
21.4984 | -106.0329
21.5023 | -106.036
21.5061 | -106.0389
21.51 | -106.0425
21.5138 | -106.0425
21.5176 | -106.0421
21.5215 | -106.0423
21.5253 | -106.0448
21.5292 | -106.0489
21.533 | -106.0499
21.5368 | -106.0506
21.5407 | -106.0507
21.5445 | -106.052
21.5484 | -106.0532
21.5522 | -106.053
21.556 | -106.0528
21.5599 | -106.0513
21.5637 | -106.0509
21.5676 | -106.0504
21.5714 | -106.05
21.5752 | -106.0474
21.5791 | -106.0445
21.5829 | -106.0426
21.5868 | -106.0409
21.5906 | -106.04
21.5944 | -106.0397
21.5983 | -106.0387
21.6021 | -106.0392
21.606 | -106.0376
21.6098 | -106.0373
21.6136 | -106.0367
21.6175 | -106.038
21.6213 | -106.0408
21.6252 | -106.0426
21.629 | -106.0437
21.6328 | -106.0462
21.6367 | -106.0508
21.6405 | -106.0552
21.6443 | -106.0588
21.6482 | -106.0637
21.652 | -106.0691
21.6559 | -106.0718
21.6597 | -106.0758
21.6635 | -106.081
21.6674 | -106.0844
21.6712 | -106.0884
21.6751 | -106.0947
21.6789 | -106.1017
21.6827 | -106.1084
21.6866 | -106.1128
21.6904 | -106.1185
21.6943 | -106.1246
21.6981 | -106.1296
21.7019 | -106.1334
21.7058 | -106.1372
21.7096 | -106.1391
21.7135 | -106.1409
21.7173 | -106.1438
21.7211 | -106.1454
21.725 | -106.1452
21.7288 | -106.1459
21.7327 | -106.1431
21.7365 | -106.1413
21.7403 | -106.1416
21.7442 | -106.1404
21.748 | -106.1375
21.7519 | -106.1315
21.7557 | -106.1246
21.7595 | -106.1201
21.7634 | -106.1164
21.7672 | -106.1094
21.7711 | -106.1023
21.7749 | -106.0946
21.7787 | -106.0879
21.7826 | -106.078
21.7864 | -106.0684
21.7903 | -106.0584
21.7941 | -106.0484
21.7979 | -106.0384
21.8018 | -106.0277
21.8056 | -106.0164
21.8095 | -106.0018
21.8133 | -105.9896
21.8171 | -105.9786
21.821 | -105.9677
21.8248 | -105.9556
21.8287 | -105.9455
21.8325 | -105.9343
21.8363 | -105.9224
21.8402 | -105.9111
21.844 | -105.9013
21.8479 | -105.8926
21.8517 | -105.8843
21.8555 | -105.8752
21.8594 | -105.8681
21.8632 | -105.858
21.8671 | -105.8496
21.8709 | -105.8427
21.8747 | -105.833
21.8786 | -105.8253
21.8824 | -105.8208
21.8862 | -105.8155
21.8901 | -105.8122
21.8939 | -105.8069
21.8978 | -105.8011
21.9016 | -105.799
21.9054 | -105.7971
21.9093 | -105.7972
21.9131 | -105.7963
21.917 | -105.794
21.9208 | -105.7937
21.9246 | -105.7947
21.9285 | -105.795
21.9323 | -105.7937
21.9362 | -105.7939
21.94 | -105.7961
21.9438 | -105.7976
21.9477 | -105.7979
21.9515 | -105.7993
21.9554 | -105.8024
21.9592 | -105.8056
21.963 | -105.811
21.9669 | -105.8173
21.9707 | -105.8214
21.9746 | -105.8264
21.9784 | -105.8321
21.9822 | -105.8378
21.9861 | -105.8425
21.9899 | -105.8482
21.9938 | -105.8534
21.9976 | -105.8591
22.0014 | -105.8638
22.0053 | -105.8695
22.0091 | -105.8755
22.013 | -105.882
22.0168 | -105.8892
22.0206 | -105.8961
22.0245 | -105.9032
22.0283 | -105.9083
22.0322 | -105.9139
22.036 | -105.921
22.0398 | -105.9276
22.0437 | -105.9342
22.0475 | -105.9403
22.0514 | -105.946
22.0552 | -105.952
22.059 | -105.9574
22.0629 | -105.9644
22.0667 | -105.9711
22.0706 | -105.9754
22.0744 | -105.9794
22.0782 | -105.9845
22.0821 | -105.9903
22.0859 | -105.9954
22.0898 | -105.9998
22.0936 | -106.0044
22.0974 | -106.0077
22.1013 | -106.0126
22.1051 | -106.0177
22.109 | -106.0227
22.1128 | -106.0298
22.1166 | -106.0369
22.1205 | -106.0453
22.1243 | -106.0518
22.1282 | -106.0581
22.132 | -106.0653
22.1358 | -106.0723
22.1397 | -106.0782
22.1435 | -106.083
22.1473 | -106.0874
22.1512 | -106.093
22.155 | -106.0976
22.1589 | -106.1012
22.1627 | -106.1053
22.1665 | -106.1104
22.1704 | -106.1152
22.1742 | -106.1197
22.1781 | -106.126
22.1819 | -106.132
22.1857 | -106.1361
22.1896 | -106.142
22.1934 | -106.1463
22.1973 | -106.1478
22.2011 | -106.1496
22.2049 | -106.1495
22.2088 | -106.1506
22.2126 | -106.153
22.2165 | -106.1512
22.2203 | -106.1485
22.2241 | -106.1453
22.228 | -106.1449
22.2318 | -106.1454
22.2357 | -106.1434
22.2395 | -106.1397
22.2433 | -106.1361
22.2472 | -106.1338
22.251 | -106.1291
22.2549 | -106.1243
22.2587 | -106.1185
22.2625 | -106.1152
22.2664 | -106.1122
22.2702 | -106.1064
22.2741 | -106.0983
22.2779 | -106.0904
22.2817 | -106.0843
22.2856 | -106.0767
22.2894 | -106.0682
22.2933 | -106.0624
22.2971 | -106.056
22.3009 | -106.0499
22.3048 | -106.0455
22.3086 | -106.0399
22.3125 | -106.0339
22.3163 | -106.0277
22.3201 | -106.022
22.324 | -106.0164
22.3278 | -106.0127
22.3317 | -106.0101
22.3355 | -106.0071
22.3393 | -106.0021
22.3432 | -105.9974
22.347 | -105.9959
22.3509 | -105.992
22.3547 | -105.9885
22.3585 | -105.9871
22.3624 | -105.9857
22.3662 | -105.9855
22.37 | -105.9848
22.3739 | -105.9846
22.3777 | -105.9849
22.3816 | -105.9842
22.3854 | -105.9817
22.3892 | -105.9813
22.3931 | -105.9775
22.3969 | -105.9714
22.4008 | -105.9683
22.4046 | -105.9669
22.4084 | -105.9662
22.4123 | -105.9657
22.4161 | -105.9643
22.42 | -105.9587
22.4238 | -105.9542
22.4276 | -105.9502
22.4315 | -105.9453
22.4353 | -105.939
22.4392 | -105.9334
22.443 | -105.927
22.4468 | -105.9185
22.4507 | -105.9121
22.4545 | -105.9032
22.4584 | -105.8973
22.4622 | -105.8882
22.466 | -105.8767
22.4699 | -105.8658
22.4737 | -105.8551
22.4776 | -105.8468
22.4814 | -105.8348
22.4852 | -105.8235
22.4891 | -105.8148
22.4929 | -105.8045
22.4968 | -105.7934
22.5006 | -105.7825
22.5044 | -105.7686
22.5083 | -105.7561
22.5121 | -105.7455
22.516 | -105.7361
22.5198 | -105.7268
22.5236 | -105.719
22.5275 | -105.7116
22.5313 | -105.7036
22.5352 | -105.6979
22.539 | -105.6938
22.5428 | -105.6893
22.5467 | -105.6851
22.5505 | -105.682
22.5544 | -105.6798
22.5582 | -105.676
22.562 | -105.6726
22.5659 | -105.6702
22.5697 | -105.6701
22.5736 | -105.6715
22.5774 | -105.676
22.5812 | -105.6786
22.5851 | -105.6813
22.5889 | -105.6852
22.5928 | -105.6927
22.5966 | -105.6995
22.6004 | -105.7046
22.6043 | -105.7122
22.6081 | -105.7211
22.612 | -105.7273
22.6158 | -105.7352
22.6196 | -105.7454
22.6235 | -105.7547
22.6273 | -105.7647
22.6312 | -105.777
22.635 | -105.7896
22.6388 | -105.8
22.6427 | -105.8107
22.6465 | -105.8211
22.6503 | -105.8317
22.6542 | -105.8436
22.658 | -105.857
22.6619 | -105.8711
22.6657 | -105.8829
22.6695 | -105.8971
22.6734 | -105.9102
22.6772 | -105.9217
22.6811 | -105.9334
22.6849 | -105.9432
22.6887 | -105.9542
22.6926 | -105.9644
22.6964 | -105.9753
22.7003 | -105.9868
22.7041 | -105.9996
22.7079 | -106.0104
22.7118 | -106.0129
22.7156 | -106.0237
22.7195 | -106.0349
22.7233 | -106.042
22.7271 | -106.0472
22.731 | -106.0514
22.7348 | -106.0559
22.7387 | -106.0605
22.7425 | -106.0685
22.7463 | -106.0749
22.7502 | -106.0798
22.754 | -106.0836
22.7579 | -106.0881
22.7617 | -106.093
22.7655 | -106.0979
22.7694 | -106.1052
22.7732 | -106.1102
22.7771 | -106.116
22.7809 | -106.1191
22.7847 | -106.1241
22.7886 | -106.1309
22.7924 | -106.1379
22.7963 | -106.1424
22.8001 | -106.1463
22.8039 | -106.1525
22.8078 | -106.1559
22.8116 | -106.1605
22.8155 | -106.1653
22.8193 | -106.1705
22.8231 | -106.1772
22.827 | -106.1859
22.8308 | -106.1938
22.8347 | -106.1981
22.8385 | -106.2013
22.8423 | -106.2059
22.8462 | -106.2124
22.85 | -106.2196
22.8539 | -106.2255
22.8577 | -106.2305
22.8615 | -106.2358
22.8654 | -106.2425
22.8692 | -106.2491
22.873 | -106.2548
22.8769 | -106.2605
22.8807 | -106.2688
22.8846 | -106.278
22.8884 | -106.2859
22.8922 | -106.2914
22.8961 | -106.2948
22.8999 | -106.2994
22.9038 | -106.3054
22.9076 | -106.3111
22.9114 | -106.3161
22.9153 | -106.3205
22.9191 | -106.3246
22.923 | -106.329
22.9268 | -106.3344
22.9306 | -106.3358
22.9345 | -106.3365
22.9383 | -106.3369
22.9422 | -106.3374
22.946 | -106.3384
22.9498 | -106.3407
22.9537 | -106.3407
22.9575 | -106.3418
22.9614 | -106.3444
22.9652 | -106.3448
22.969 | -106.3439
22.9729 | -106.3424
22.9767 | -106.3402
22.9806 | -106.3384
22.9844 | -106.3368
22.9882 | -106.3354
22.9921 | -106.332
22.9959 | -106.3285
22.9998 | -106.3246
23.0036 | -106.3206
23.0074 | -106.3195
23.0113 | -106.3174
23.0151 | -106.3156
23.019 | -106.3128
23.0228 | -106.312
23.0266 | -106.3094
23.0305 | -106.3064
23.0343 | -106.3069
23.0382 | -106.3066
23.042 | -106.3033
23.0458 | -106.3005
23.0497 | -106.2992
23.0535 | -106.2994
23.0574 | -106.2998
23.0612 | -106.2991
23.065 | -106.297
23.0689 | -106.2947
23.0727 | -106.2933
23.0766 | -106.2925
23.0804 | -106.2925
23.0842 | -106.2937
23.0881 | -106.293
23.0919 | -106.2927
23.0958 | -106.2931
23.0996 | -106.2933
23.1034 | -106.2939
23.1073 | -106.2957
23.1111 | -106.2987
23.115 | -106.2991
23.1188 | -106.2976
23.1226 | -106.2975
23.1265 | -106.2954
23.1303 | -106.2926
23.1341 | -106.2911
23.138 | -106.29
23.1418 | -106.2891
23.1457 | -106.2875
23.1495 | -106.286
23.1533 | -106.2866
23.1572 | -106.284
23.161 | -106.2839
23.1649 | -106.2861
23.1687 | -106.2853
23.1725 | -106.2806
23.1764 | -106.2777
23.1802 | -106.2768
23.1841 | -106.2757
23.1879 | -106.2741
23.1917 | -106.2703
23.1956 | -106.2684
23.1994 | -106.2666
23.2033 | -106.2641
23.2071 | -106.2616
23.2109 | -106.2592
23.2148 | -106.2564
23.2186 | -106.2527
23.2225 | -106.2512
23.2263 | -106.2493
23.2301 | -106.2453
23.234 | -106.2412
23.2378 | -106.238
23.2417 | -106.2333
23.2455 | -106.2281
23.2493 | -106.2235
23.2532 | -106.2141
23.257 | -106.2076
23.2609 | -106.2029
23.2647 | -106.1961
23.2685 | -106.1894
23.2724 | -106.1806
23.2762 | -106.1722
23.2801 | -106.1623
23.2839 | -106.1527
23.2877 | -106.1435
23.2916 | -106.133
23.2954 | -106.1244
23.2993 | -106.118
23.3031 | -106.1113
23.3069 | -106.1047
23.3108 | -106.0973
23.3146 | -106.0901
23.3185 | -106.083
23.3223 | -106.0759
23.3261 | -106.0697
23.33 | -106.0663
23.3338 | -106.064
23.3377 | -106.0584
23.3415 | -106.0523
23.3453 | -106.0481
23.3492 | -106.0465
23.353 | -106.0441
23.3569 | -106.0432
23.3607 | -106.0416
23.3645 | -106.0373
23.3684 | -106.0356
23.3722 | -106.0355
23.376 | -106.0364
23.3799 | -106.0364
23.3837 | -106.0359
23.3876 | -106.0365
23.3914 | -106.0371
23.3952 | -106.0379
23.3991 | -106.0385
23.4029 | -106.0393
23.4068 | -106.0392
23.4106 | -106.0406
23.4144 | -106.0414
23.4183 | -106.0421
23.4221 | -106.0425
23.426 | -106.0423
23.4298 | -106.044
23.4336 | -106.0462
23.4375 | -106.0488
23.4413 | -106.0486
23.4452 | -106.0486
23.449 | -106.0502
23.4528 | -106.0508
23.4567 | -106.052
23.4605 | -106.0538
23.4644 | -106.0549
23.4682 | -106.0532
23.472 | -106.0504
23.4759 | -106.0484
23.4797 | -106.0464
23.4836 | -106.0432
23.4874 | -106.0428
23.4912 | -106.0405
23.4951 | -106.038
23.4989 | -106.0358
23.5028 | -106.0338
23.5066 | -106.0308
23.5104 | -106.0276
23.5143 | -106.0248
23.5181 | -106.0202
23.522 | -106.0174
23.5258 | -106.0144
23.5296 | -106.012
23.5335 | -106.0108
23.5373 | -106.0097
23.5412 | -106.0077
23.545 | -106.0064
23.5488 | -106.007
23.5527 | -106.0062
23.5565 | -106.0039
23.5604 | -106.004
23.5642 | -106.0047
23.568 | -106.0059
23.5719 | -106.0057
23.5757 | -106.0067
23.5796 | -106.0072
23.5834 | -106.0096
23.5872 | -106.0111
23.5911 | -106.0107
23.5949 | -106.0097
23.5988 | -106.0093
23.6026 | -106.0096
23.6064 | -106.0103
23.6103 | -106.0104
23.6141 | -106.0101
23.618 | -106.0109
23.6218 | -106.0112
23.6256 | -106.0096
23.6295 | -106.0086
23.6333 | -106.0062
23.6371 | -106.0028
23.641 | -105.9985
23.6448 | -105.9945
23.6487 | -105.99
23.6525 | -105.9845
23.6563 | -105.9781
23.6602 | -105.9698
23.664 | -105.9605
23.6679 | -105.9484
23.6717 | -105.939
23.6755 | -105.9318
23.6794 | -105.9225
23.6832 | -105.9105
23.6871 | -105.899
23.6909 | -105.8864
23.6947 | -105.873
23.6986 | -105.8617
23.7024 | -105.8481
23.7063 | -105.8329
23.7101 | -105.8163
23.7139 | -105.7999
23.7178 | -105.7848
23.7216 | -105.7689
23.7255 | -105.7539
23.7293 | -105.7383
23.7331 | -105.7225
23.737 | -105.7072
23.7408 | -105.6943
23.7447 | -105.6816
23.7485 | -105.6684
23.7523 | -105.6561
23.7562 | -105.6446
23.76 | -105.6331
23.7639 | -105.6209
23.7677 | -105.6085
23.7715 | -105.5968
23.7754 | -105.5876
23.7792 | -105.5799
23.7831 | -105.5722
23.7869 | -105.5641
23.7907 | -105.5557
23.7946 | -105.549
23.7984 | -105.5421
23.8023 | -105.5372
23.8061 | -105.5333
23.8099 | -105.531
23.8138 | -105.5276
23.8176 | -105.5316
23.8215 | -105.5309
23.8253 | -105.5299
23.8291 | -105.5296
23.833 | -105.5297
23.8368 | -105.5302
23.8407 | -105.5317
23.8445 | -105.5331
23.8483 | -105.5333
23.8522 | -105.5335
23.856 | -105.5376
23.8599 | -105.5424
23.8637 | -105.5451
23.8675 | -105.5492
23.8714 | -105.5531
23.8752 | -105.5561
23.879 | -105.5594
23.8829 | -105.5633
23.8867 | -105.5666
23.8906 | -105.5715
23.8944 | -105.5765
23.8982 | -105.5811
23.9021 | -105.5845
23.9059 | -105.5865
23.9098 | -105.5915
23.9136 | -105.596
23.9174 | -105.5982
23.9213 | -105.6
23.9251 | -105.6028
23.929 | -105.6058
23.9328 | -105.6053
23.9366 | -105.6039
23.9405 | -105.6034
23.9443 | -105.602
23.9482 | -105.6002
23.952 | -105.5993
23.9558 | -105.5971
23.9597 | -105.5935
23.9635 | -105.5914
23.9674 | -105.5882
23.9712 | -105.5842
23.975 | -105.5801
23.9789 | -105.5773
23.9827 | -105.5748
23.9866 | -105.5727
23.9904 | -105.5693
23.9942 | -105.5663
23.9981 | -105.566
24.0019 | -105.5636
24.0058 | -105.5617
24.0096 | -105.5585
24.0134 | -105.5547
24.0173 | -105.5517
24.0211 | -105.5508
24.025 | -105.5503
24.0288 | -105.5492
24.0326 | -105.5504
24.0365 | -105.5499
24.0403 | -105.5502
24.0442 | -105.551
24.048 | -105.552
24.0518 | -105.5547
24.0557 | -105.5568
24.0595 | -105.5587
24.0634 | -105.5605
24.0672 | -105.5638
24.071 | -105.5682
24.0749 | -105.5708
24.0787 | -105.5729
24.0826 | -105.5768
24.0864 | -105.5816
24.0902 | -105.5857
24.0941 | -105.5885
24.0979 | -105.592
24.1018 | -105.5966
24.1056 | -105.6013
24.1094 | -105.6072
24.1133 | -105.6125
24.1171 | -105.6155
24.121 | -105.6193
24.1248 | -105.6242
24.1286 | -105.6294
24.1325 | -105.6338
24.1363 | -105.6398
24.1401 | -105.6426
24.144 | -105.6447
24.1478 | -105.6435
24.1517 | -105.6433
24.1555 | -105.6433
24.1593 | -105.6423
24.1632 | -105.6412
24.167 | -105.6399
24.1709 | -105.6388
24.1747 | -105.6378
24.1785 | -105.6351
24.1824 | -105.6323
24.1862 | -105.6283
24.1901 | -105.6248
24.1939 | -105.6222
24.1977 | -105.6176
24.2016 | -105.6108
24.2054 | -105.605
24.2093 | -105.5992
24.2131 | -105.5921
24.2169 | -105.5851
24.2208 | -105.5791
24.2246 | -105.5715
24.2285 | -105.5648
24.2323 | -105.5567
24.2361 | -105.5474
24.24 | -105.5385
24.2438 | -105.5293
24.2477 | -105.5219
24.2515 | -105.5143
24.2553 | -105.5057
24.2592 | -105.4996
24.263 | -105.4874
24.2669 | -105.4704
24.2707 | -105.4562
24.2745 | -105.4361
24.2784 | -105.4039
24.2822 | -105.3553
24.2861 | -105.3155
24.2899 | -105.2849
24.2937 | -105.2575
24.2976 | -105.2286
24.3014 | -105.198
24.3053 | -105.1678
24.3091 | -105.1363
24.3129 | -105.1028
24.3168 | -105.0692
24.3206 | -105.0335
24.3245 | -104.9946
24.3283 | -104.9538
24.3321 | -104.9128
24.336 | -104.8695
24.3398 | -104.8238
24.3437 | -104.7778
24.3475 | -104.7305
24.3513 | -104.6811
24.3552 | -104.631
24.359 | -104.5817
24.3629 | -104.5282
24.3667 | -104.4747
24.3705 | -104.4197
24.3744 | -104.3709
24.3782 | -104.3233
24.382 | -104.2758
24.3859 | -104.2287
24.3897 | -104.1823
24.3936 | -104.1362
24.3974 | -104.091
24.4012 | -104.0493
24.4051 | -104.011
24.4089 | -103.9745
24.4128 | -103.9398
24.4166 | -103.9072
24.4204 | -103.878
24.4243 | -103.8497
24.4281 | -103.8242
24.432 | -103.8033
24.4358 | -103.785
24.4396 | -103.7695
24.4435 | -103.7583
24.4473 | -103.7482
24.4512 | -103.7411
24.455 | -103.7373
24.4588 | -103.7375
24.4627 | -103.7416
24.4665 | -103.7476
24.4704 | -103.7578
24.4742 | -103.7721
24.478 | -103.7879
24.4819 | -103.8055
24.4857 | -103.8271
24.4896 | -103.8524
24.4934 | -103.8806
24.4972 | -103.9107
24.5011 | -103.9418
24.5049 | -103.9748
24.5088 | -104.0095
24.5126 | -104.0455
24.5164 | -104.082
24.5203 | -104.1195
24.5241 | -104.1599
24.528 | -104.2001
24.5318 | -104.2394
24.5356 | -104.2787
24.5395 | -104.3181
24.5433 | -104.3579
24.5472 | -104.397
24.551 | -104.4354
24.5548 | -104.4731
24.5587 | -104.5099
24.5625 | -104.545
24.5664 | -104.5741
24.5702 | -104.5872
24.574 | -104.5878
24.5779 | -104.5594
24.5817 | -104.5146
24.5856 | -104.4421
24.5894 | -104.3715
24.5932 | -104.3099
24.5971 | -104.2485
24.6009 | -104.1993
24.6048 | -104.1469
24.6086 | -104.0904
24.6124 | -104.0284
24.6163 | -103.9567
24.6201 | -103.8813
24.624 | -103.8048
24.6278 | -103.7287
24.6316 | -103.6388
24.6355 | -103.5361
24.6393 | -103.4249
24.6431 | -103.3103
24.647 | -103.1945
24.6508 | -103.0753
24.6547 | -102.9534
24.6585 | -102.8304
24.6623 | -102.7066
24.6662 | -102.5828
24.67 | -102.4601
24.6739 | -102.3386
24.6777 | -102.2187
24.6815 | -102.1
24.6854 | -101.9838
24.6892 | -101.871
24.6931 | -101.7611
24.6969 | -101.655
24.7007 | -101.5533
24.7046 | -101.4555
24.7084 | -101.3637
24.7123 | -101.277
24.7161 | -101.1954
24.7199 | -101.1192
24.7238 | -101.0498
24.7276 | -100.9866
24.7315 | -100.9308
24.7353 | -100.8818
24.7391 | -100.8394
24.743 | -100.8039
24.7468 | -100.7752
24.7507 | -100.753
24.7545 | -100.7379
24.7583 | -100.731
24.7622 | -100.7317
24.766 | -100.7398
24.7699 | -100.7551
24.7737 | -100.7784
24.7775 | -100.8084
24.7814 | -100.8449
24.7852 | -100.8882
24.7891 | -100.9388
24.7929 | -100.9956
24.7967 | -101.0597
24.8006 | -101.13
24.8044 | -101.2048
24.8083 | -101.286
24.8121 | -101.373
24.8159 | -101.4646
24.8198 | -101.5599
24.8236 | -101.6593
24.8275 | -101.7642
24.8313 | -101.8721
24.8351 | -101.982
24.839 | -102.0938
24.8428 | -102.2077
24.8467 | -102.3235
24.8505 | -102.4407
24.8543 | -102.559
24.8582 | -102.6768
24.862 | -102.7949
24.8659 | -102.9134
24.8697 | -103.0294
24.8735 | -103.1417
24.8774 | -103.2525
24.8812 | -103.3624
24.885 | -103.4699
24.8889 | -103.5729
24.8927 | -103.671
24.8966 | -103.7656
24.9004 | -103.8565
24.9042 | -103.9427
24.9081 | -104.0259
24.9119 | -104.1043
24.9158 | -104.1804
24.9196 | -104.2645
24.9234 | -104.3495
24.9273 | -104.3657
24.9311 | -104.4612
24.935 | -104.5572
24.9388 | -104.662
24.9426 | -104.7451
24.9465 | -104.7983
24.9503 | -104.8368
24.9542 | -104.8508
24.958 | -104.8602
24.9618 | -104.8688
24.9657 | -104.8776
24.9695 | -104.886
24.9734 | -104.8951
24.9772 | -104.9038
24.981 | -104.9126
24.9849 | -104.9188
24.9887 | -104.9218
24.9926 | -104.9262
24.9964 | -104.9308
25.0002 | -104.9328
25.0041 | -104.9356
25.0079 | -104.9384
25.0118 | -104.9416
25.0156 | -104.9459
25.0194 | -104.9501
25.0233 | -104.9541
25.0271 | -104.9587
25.031 | -104.9626
25.0348 | -104.9674
25.0386 | -104.9717
25.0425 | -104.9762
25.0463 | -104.9822
25.0502 | -104.9887
25.054 | -104.9923
25.0578 | -104.9954
25.0617 | -104.9996
25.0655 | -105.0045
25.0694 | -105.0086
25.0732 | -105.0138
25.077 | -105.0168
25.0809 | -105.0187
25.0847 | -105.0207
25.0886 | -105.0235
25.0924 | -105.0263
25.0962 | -105.0292
25.1001 | -105.0329
25.1039 | -105.0356
25.1078 | -105.0379
25.1116 | -105.0373
25.1154 | -105.0369
25.1193 | -105.0358
25.1231 | -105.0375
25.127 | -105.0377
25.1308 | -105.0383
25.1346 | -105.0382
25.1385 | -105.0384
25.1423 | -105.0386
25.1461 | -105.0367
25.15 | -105.0333
25.1538 | -105.0274
25.1577 | -105.0249
25.1615 | -105.021
25.1653 | -105.0166
25.1692 | -105.0129
25.173 | -105.0088
25.1769 | -105.0062
25.1807 | -105.0025
25.1845 | -104.9989
25.1884 | -104.9963
25.1922 | -104.9927
25.1961 | -104.9894
25.1999 | -104.9856
25.2037 | -104.9816
25.2076 | -104.9762
25.2114 | -104.9701
25.2153 | -104.9652
25.2191 | -104.9585
25.2229 | -104.9528
25.2268 | -104.9494
25.2306 | -104.9456
25.2345 | -104.9395
25.2383 | -104.9335
25.2421 | -104.9274
25.246 | -104.9207
25.2498 | -104.9142
25.2537 | -104.9072
25.2575 | -104.9001
25.2613 | -104.8921
25.2652 | -104.8844
25.269 | -104.8769
25.2729 | -104.8705
25.2767 | -104.8631
25.2805 | -104.8572
25.2844 | -104.8525
25.2882 | -104.8459
25.2921 | -104.8386
25.2959 | -104.8326
25.2997 | -104.8264
25.3036 | -104.8213
25.3074 | -104.8156
25.3113 | -104.8101
25.3151 | -104.8034
25.3189 | -104.797
25.3228 | -104.7898
25.3266 | -104.7836
25.3305 | -104.7787
25.3343 | -104.7743
25.3381 | -104.7684
25.342 | -104.762
25.3458 | -104.7584
25.3497 | -104.755
25.3535 | -104.751
25.3573 | -104.7478
25.3612 | -104.7432
25.365 | -104.7407
25.3688 | -104.7378
25.3727 | -104.7339
25.3765 | -104.7302
25.3804 | -104.7253
25.3842 | -104.7196
25.388 | -104.7158
25.3919 | -104.7134
25.3957 | -104.7104
25.3996 | -104.7052
25.4034 | -104.7015
25.4072 | -104.6984
25.4111 | -104.6966
25.4149 | -104.6958
25.4188 | -104.6929
25.4226 | -104.69
25.4264 | -104.6889
25.4303 | -104.6865
25.4341 | -104.6854
25.438 | -104.6842
25.4418 | -104.6827
25.4456 | -104.6815
25.4495 | -104.6798
25.4533 | -104.6784
25.4572 | -104.677
25.461 | -104.677
25.4648 | -104.6766
25.4687 | -104.6762
25.4725 | -104.6759
25.4764 | -104.6761
25.4802 | -104.6767
25.484 | -104.675
25.4879 | -104.6727
25.4917 | -104.6718
25.4956 | -104.6725
25.4994 | -104.673
25.5032 | -104.6731
25.5071 | -104.6745
25.5109 | -104.6766
25.5148 | -104.6802
25.5186 | -104.6812
25.5224 | -104.6819
25.5263 | -104.6834
25.5301 | -104.6843
25.534 | -104.6861
25.5378 | -104.6854
25.5416 | -104.6846
25.5455 | -104.6862
25.5493 | -104.6893
25.5532 | -104.6929
25.557 | -104.6963
25.5608 | -104.6996
25.5647 | -104.7037
25.5685 | -104.7087
25.5724 | -104.7143
25.5762 | -104.719
25.58 | -104.7235
25.5839 | -104.7277
25.5877 | -104.7321
25.5916 | -104.7378
25.5954 | -104.7446
25.5992 | -104.7512
25.6031 | -104.7583
25.6069 | -104.7652
25.6108 | -104.7733
25.6146 | -104.7821
25.6184 | -104.7909
25.6223 | -104.7993
25.6261 | -104.8086
25.63 | -104.8199
25.6338 | -104.8288
25.6376 | -104.8385
25.6415 | -104.8469
25.6453 | -104.8547
25.6491 | -104.8644
25.653 | -104.8752
25.6568 | -104.8829
25.6607 | -104.8915
25.6645 | -104.9008
25.6683 | -104.9098
25.6722 | -104.9197
25.676 | -104.9315
25.6799 | -104.9444
25.6837 | -104.9565
25.6875 | -104.9686
25.6914 | -104.9792
25.6952 | -104.9871
25.6991 | -104.9963
25.7029 | -105.0074
25.7067 | -105.0159
25.7106 | -105.0239
25.7144 | -105.0346
25.7183 | -105.0438
25.7221 | -105.0534
25.7259 | -105.0616
25.7298 | -105.072
25.7336 | -105.0806
25.7375 | -105.0902
25.7413 | -105.1
25.7451 | -105.1093
25.749 | -105.1177
25.7528 | -105.1241
25.7567 | -105.1302
25.7605 | -105.1379
25.7643 | -105.1438
25.7682 | -105.1497
25.772 | -105.1547
25.7759 | -105.1581
25.7797 | -105.1626
25.7835 | -105.1664
25.7874 | -105.1673
25.7912 | -105.1677
25.7951 | -105.1684
25.7989 | -105.1685
25.8027 | -105.1687
25.8066 | -105.1665
25.8104 | -105.164
25.8143 | -105.1629
25.8181 | -105.1619
25.8219 | -105.159
25.8258 | -105.1566
25.8296 | -105.1546
25.8335 | -105.1505
25.8373 | -105.145
25.8411 | -105.1414
25.845 | -105.1362
25.8488 | -105.1307
25.8527 | -105.1271
25.8565 | -105.1229
25.8603 | -105.117
25.8642 | -105.1114
25.868 | -105.1076
25.8718 | -105.1032
25.8757 | -105.098
25.8795 | -105.0919
25.8834 | -105.0874
25.8872 | -105.0825
25.891 | -105.0768
25.8949 | -105.073
25.8987 | -105.0674
25.9026 | -105.065
25.9064 | -105.0643
25.9102 | -105.0612
25.9141 | -105.0568
25.9179 | -105.0533
25.9218 | -105.0502
25.9256 | -105.0476
25.9294 | -105.0446
25.9333 | -105.0419
25.9371 | -105.0391
25.941 | -105.0351
25.9448 | -105.0327
25.9486 | -105.0319
25.9525 | -105.0301
25.9563 | -105.0275
25.9602 | -105.0247
25.964 | -105.0219
25.9678 | -105.0191
25.9717 | -105.0167
25.9755 | -105.0142
25.9794 | -105.0116
25.9832 | -105.0081
25.987 | -105.0063
25.9909 | -105.0038
25.9947 | -104.9987
25.9986 | -104.9963
26.0024 | -104.9944
26.0062 | -104.9929
26.0101 | -104.9912
26.0139 | -104.9896
26.0178 | -104.9879
26.0216 | -104.9875
26.0254 | -104.9872
26.0293 | -104.9871
26.0331 | -104.9889
26.037 | -104.9894
26.0408 | -104.99
26.0446 | -104.9896
26.0485 | -104.9888
26.0523 | -104.988
26.0562 | -104.9856
26.06 | -104.9838
26.0638 | -104.9867
26.0677 | -104.989
26.0715 | -104.9889
26.0754 | -104.9902
26.0792 | -104.993
26.083 | -104.9927
26.0869 | -104.9932
26.0907 | -104.9961
26.0946 | -104.9989
26.0984 | -105.0002
26.1022 | -105.0027
26.1061 | -105.0038
26.1099 | -105.0062
26.1138 | -105.0079
26.1176 | -105.0095
26.1214 | -105.0114
26.1253 | -105.0136
26.1291 | -105.0162
26.1329 | -105.0181
26.1368 | -105.02
26.1406 | -105.0224
26.1445 | -105.026
26.1483 | -105.0293
26.1521 | -105.0301
26.156 | -105.0323
26.1598 | -105.0366
26.1637 | -105.0396
26.1675 | -105.0399
26.1713 | -105.0392
26.1752 | -105.0413
26.179 | -105.0401
26.1829 | -105.0384
26.1867 | -105.0358
26.1905 | -105.0344
26.1944 | -105.0333
26.1982 | -105.0326
26.2021 | -105.0309
26.2059 | -105.0287
26.2097 | -105.0271
26.2136 | -105.0248
26.2174 | -105.0236
26.2213 | -105.0222
26.2251 | -105.0188
26.2289 | -105.0147
26.2328 | -105.0122
26.2366 | -105.0084
26.2405 | -105.0028
26.2443 | -104.9966
26.2481 | -104.992
26.252 | -104.9867
26.2558 | -104.9816
26.2597 | -104.9778
26.2635 | -104.9735
26.2673 | -104.9697
26.2712 | -104.9636
26.275 | -104.9559
26.2789 | -104.95
26.2827 | -104.9456
26.2865 | -104.9416
26.2904 | -104.9363
26.2942 | -104.9319
26.2981 | -104.9284
26.3019 | -104.9251
26.3057 | -104.9227
26.3096 | -104.9206
26.3134 | -104.9185
26.3173 | -104.916
26.3211 | -104.9146
26.3249 | -104.9132
26.3288 | -104.9125
26.3326 | -104.9137
26.3365 | -104.9155
26.3403 | -104.9175
26.3441 | -104.9185
26.348 | -104.9198
26.3518 | -104.924
26.3557 | -104.9278
26.3595 | -104.9336
26.3633 | -104.9415
26.3672 | -104.9473
26.371 | -104.9533
26.3748 | -104.9596
26.3787 | -104.9684
26.3825 | -104.9773
26.3864 | -104.9866
26.3902 | -104.9967
26.394 | -105.0085
26.3979 | -105.0198
26.4017 | -105.0295
26.4056 | -105.0413
26.4094 | -105.0534
26.4132 | -105.0665
26.4171 | -105.0803
26.4209 | -105.0938
26.4248 | -105.1094
26.4286 | -105.1266
26.4324 | -105.1431
26.4363 | -105.1574
26.4401 | -105.1711
26.444 | -105.1855
26.4478 | -105.2004
26.4516 | -105.2138
26.4555 | -105.2278
26.4593 | -105.243
26.4632 | -105.256
26.467 | -105.2685
26.4708 | -105.2811
26.4747 | -105.2926
26.4785 | -105.3038
26.4824 | -105.3166
26.4862 | -105.3298
26.49 | -105.3428
26.4939 | -105.3552
26.4977 | -105.3677
26.5016 | -105.3779
26.5054 | -105.3873
26.5092 | -105.3959
26.5131 | -105.406
26.5169 | -105.415
26.5208 | -105.4228
26.5246 | -105.4346
26.5284 | -105.4461
26.5323 | -105.4544
26.5361 | -105.4612
26.54 | -105.4688
26.5438 | -105.477
26.5476 | -105.4848
26.5515 | -105.4957
26.5553 | -105.5033
26.5592 | -105.5099
26.563 | -105.5157
26.5668 | -105.5212
26.5707 | -105.5266
26.5745 | -105.5336
26.5784 | -105.5406
26.5822 | -105.545
26.586 | -105.5446
26.5899 | -105.5459
26.5937 | -105.5473
26.5976 | -105.5476
26.6014 | -105.5477
26.6052 | -105.548
26.6091 | -105.5484
26.6129 | -105.5476
26.6168 | -105.5475
26.6206 | -105.549
26.6244 | -105.5499
26.6283 | -105.5503
26.6321 | -105.5494
26.6359 | -105.5473
26.6398 | -105.5451
26.6436 | -105.5423
26.6475 | -105.5401
26.6513 | -105.5387
26.6551 | -105.5377
26.659 | -105.5367
26.6628 | -105.5327
26.6667 | -105.5297
26.6705 | -105.5263
26.6743 | -105.5227
26.6782 | -105.5184
26.682 | -105.515
26.6859 | -105.5134
26.6897 | -105.5112
26.6935 | -105.5095
26.6974 | -105.506
26.7012 | -105.5023
26.7051 | -105.5
26.7089 | -105.4969
26.7127 | -105.4947
26.7166 | -105.493
26.7204 | -105.4922
26.7243 | -105.4902
26.7281 | -105.4869
26.7319 | -105.4846
26.7358 | -105.4808
26.7396 | -105.4764
26.7435 | -105.4742
26.7473 | -105.4723
26.7511 | -105.4703
26.755 | -105.4676
26.7588 | -105.4653
26.7627 | -105.463
26.7665 | -105.4614
26.7703 | -105.4582
26.7742 | -105.456
26.778 | -105.4564
26.7819 | -105.4562
26.7857 | -105.4579
26.7895 | -105.4586
26.7934 | -105.4598
26.7972 | -105.4625
26.8011 | -105.4629
26.8049 | -105.4628
26.8087 | -105.4639
26.8126 | -105.467
26.8164 | -105.4703
26.8203 | -105.4749
26.8241 | -105.4788
26.8279 | -105.481
26.8318 | -105.484
26.8356 | -105.4868
26.8395 | -105.491
26.8433 | -105.4962
26.8471 | -105.5012
26.851 | -105.5078
26.8548 | -105.5156
26.8587 | -105.5206
26.8625 | -105.5239
26.8663 | -105.5284
26.8702 | -105.5343
26.874 | -105.5384
26.8778 | -105.5436
26.8817 | -105.5496
26.8855 | -105.5561
26.8894 | -105.5589
26.8932 | -105.5627
26.897 | -105.567
26.9009 | -105.569
26.9047 | -105.5691
26.9086 | -105.5705
26.9124 | -105.5752
26.9162 | -105.5797
26.9201 | -105.5818
26.9239 | -105.5812
26.9278 | -105.5803
26.9316 | -105.5803
26.9354 | -105.5807
26.9393 | -105.5817
26.9431 | -105.5812
26.947 | -105.5807
26.9508 | -105.5794
26.9546 | -105.5781
26.9585 | -105.5764
26.9623 | -105.5743
26.9662 | -105.5738
26.97 | -105.5725
26.9738 | -105.5683
26.9777 | -105.5658
26.9815 | -105.5649
26.9854 | -105.5641
26.9892 | -105.5623
26.993 | -105.5594
26.9969 | -105.5569
27.0007 | -105.5531
27.0046 | -105.5487
27.0084 | -105.544
27.0122 | -105.5402
27.0161 | -105.5368
27.0199 | -105.5332
27.0238 | -105.5315
27.0276 | -105.5282
27.0314 | -105.5251
27.0353 | -105.5207
27.0391 | -105.5166
27.043 | -105.5136
27.0468 | -105.5113
27.0506 | -105.5076
27.0545 | -105.5034
27.0583 | -105.5023
27.0622 | -105.4999
27.066 | -105.4962
27.0698 | -105.4933
27.0737 | -105.4899
27.0775 | -105.4881
27.0814 | -105.4872
27.0852 | -105.485
27.089 | -105.4823
27.0929 | -105.4804
27.0967 | -105.4794
27.1006 | -105.4784
27.1044 | -105.4788
27.1082 | -105.479
27.1121 | -105.4784
27.1159 | -105.4783
27.1197 | -105.4784
27.1236 | -105.4786
27.1274 | -105.4792
27.1313 | -105.4798
27.1351 | -105.4799
27.1389 | -105.4786
27.1428 | -105.4792
27.1466 | -105.4787
27.1505 | -105.4772
27.1543 | -105.4768
27.1581 | -105.4769
27.162 | -105.477
27.1658 | -105.476
27.1697 | -105.4738
27.1735 | -105.4706
27.1773 | -105.4691
27.1812 | -105.4675
27.185 | -105.4676
27.1889 | -105.4687
27.1927 | -105.4694
27.1965 | -105.468
27.2004 | -105.4667
27.2042 | -105.466
27.2081 | -105.4667
27.2119 | -105.4674
27.2157 | -105.4665
27.2196 | -105.4659
27.2234 | -105.4649
27.2273 | -105.4618
27.2311 | -105.4581
27.2349 | -105.4558
27.2388 | -105.4532
27.2426 | -105.4509
27.2465 | -105.4494
27.2503 | -105.4479
27.2541 | -105.4469
27.258 | -105.4465
27.2618 | -105.444
27.2657 | -105.4415
27.2695 | -105.4387
27.2733 | -105.4379
27.2772 | -105.4363
27.281 | -105.4348
27.2849 | -105.434
27.2887 | -105.4343
27.2925 | -105.4341
27.2964 | -105.4318
27.3002 | -105.4309
27.3041 | -105.4304
27.3079 | -105.4295
27.3117 | -105.429
27.3156 | -105.425
27.3194 | -105.4208
27.3233 | -105.4173
27.3271 | -105.4159
27.3309 | -105.4149
27.3348 | -105.4151
27.3386 | -105.4138
27.3425 | -105.4127
27.3463 | -105.4144
27.3501 | -105.416
27.354 | -105.4171
27.3578 | -105.4189
27.3617 | -105.42
27.3655 | -105.4196
27.3693 | -105.4195
27.3732 | -105.4188
27.377 | -105.4172
27.3808 | -105.4169
27.3847 | -105.4182
27.3885 | -105.4189
27.3924 | -105.4199
27.3962 | -105.4224
27.4 | -105.4241
27.4039 | -105.425
27.4077 | -105.427
27.4116 | -105.4302
27.4154 | -105.4312
27.4192 | -105.4335
27.4231 | -105.4342
27.4269 | -105.4364
27.4308 | -105.4381
27.4346 | -105.4404
27.4384 | -105.4421
27.4423 | -105.4428
27.4461 | -105.4442
27.45 | -105.4455
27.4538 | -105.4471
27.4576 | -105.4487
27.4615 | -105.4482
27.4653 | -105.4484
27.4692 | -105.4506
27.473 | -105.4528
27.4768 | -105.4533
27.4807 | -105.4544
27.4845 | -105.4546
27.4884 | -105.4544
27.4922 | -105.4551
27.496 | -105.4548
27.4999 | -105.4563
27.5037 | -105.4574
27.5076 | -105.4578
27.5114 | -105.4589
27.5152 | -105.461
27.5191 | -105.4635
27.5229 | -105.4662
27.5268 | -105.4679
27.5306 | -105.4687
27.5344 | -105.47
27.5383 | -105.4718
27.5421 | -105.473
27.546 | -105.4733
27.5498 | -105.4724
27.5536 | -105.4729
27.5575 | -105.474
27.5613 | -105.4763
27.5652 | -105.4771
27.569 | -105.4787
27.5728 | -105.4791
27.5767 | -105.4782
27.5805 | -105.4774
27.5844 | -105.4774
27.5882 | -105.4786
27.592 | -105.4783
27.5959 | -105.4764
27.5997 | -105.4755
27.6036 | -105.4752
27.6074 | -105.4739
27.6112 | -105.4703
27.6151 | -105.4685
27.6189 | -105.4671
27.6227 | -105.4646
27.6266 | -105.4606
27.6304 | -105.4568
27.6343 | -105.4548
27.6381 | -105.4518
27.6419 | -105.4473
27.6458 | -105.4435
27.6496 | -105.4418
27.6535 | -105.4406
27.6573 | -105.4377
27.6611 | -105.4354
27.665 | -105.4323
27.6688 | -105.4287
27.6727 | -105.4267
27.6765 | -105.4239
27.6803 | -105.4211
27.6842 | -105.4173
27.688 | -105.4139
27.6919 | -105.4105
27.6957 | -105.41
27.6995 | -105.4096
27.7034 | -105.4089
27.7072 | -105.4118
27.7111 | -105.412
27.7149 | -105.4117
27.7187 | -105.4091
27.7226 | -105.4069
27.7264 | -105.405
27.7303 | -105.4051
27.7341 | -105.405
27.7379 | -105.4053
27.7418 | -105.4049
27.7456 | -105.4064
27.7495 | -105.4081
27.7533 | -105.4081
27.7571 | -105.4088
27.761 | -105.4095
27.7648 | -105.4112
27.7687 | -105.4149
27.7725 | -105.4172
27.7763 | -105.4168
27.7802 | -105.4173
27.784 | -105.4185
27.7879 | -105.4183
27.7917 | -105.4182
27.7955 | -105.4193
27.7994 | -105.4197
27.8032 | -105.4198
27.8071 | -105.4187
27.8109 | -105.4177
27.8147 | -105.4175
27.8186 | -105.4177
27.8224 | -105.4178
27.8263 | -105.4176
27.8301 | -105.4171
27.8339 | -105.4163
27.8378 | -105.4146
27.8416 | -105.413
27.8455 | -105.411
27.8493 | -105.4096
27.8531 | -105.4092
27.857 | -105.4055
27.8608 | -105.399
27.8647 | -105.3943
27.8685 | -105.3918
27.8723 | -105.3897
27.8762 | -105.3895
27.88 | -105.3876
27.8838 | -105.3849
27.8877 | -105.3817
27.8915 | -105.3772
27.8954 | -105.373
27.8992 | -105.368
27.903 | -105.3637
27.9069 | -105.361
27.9107 | -105.3561
27.9146 | -105.3515
27.9184 | -105.3458
27.9222 | -105.337
27.9261 | -105.3222
27.9299 | -105.3078
27.9338 | -105.296
27.9376 | -105.2872
27.9414 | -105.2792
27.9453 | -105.27
27.9491 | -105.2603
27.953 | -105.2498
27.9568 | -105.2393
27.9606 | -105.2279
27.9645 | -105.2177
27.9683 | -105.2072
27.9722 | -105.1969
27.976 | -105.1855
27.9798 | -105.1735
27.9837 | -105.1612
27.9875 | -105.1499
27.9914 | -105.1382
27.9952 | -105.1245
27.999 | -105.1102
28.0029 | -105.0971