## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 276 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 282
Hood | 02-92B/P | 224
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1011.8 | hPa |  | PRC PRO 01
Humidity | 61.2 | %/Hr
Internal temperature of calibration bench | 24.27 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1021 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.2 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.94 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.01 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3089 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 21/08/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 24/08/2018
Technical Manager | Anthony HUE | Approval |  | 30/08/2018
CEO | Robert GLEMOT | Approval |  | 30/08/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3116 | 57.5523
0.012209 | -35.8758 | 38.3576
0.018313 | -34.8163 | 28.169
0.024417 | -34.3587 | 21.9475
0.030521 | -34.2467 | 18.1917
0.036626 | -34.1034 | 15.1341
0.04273 | -34.0061 | 12.7661
0.048834 | -33.958 | 11.2308
0.054938 | -33.8936 | 10.0892
0.061043 | -33.835 | 9.0133
0.067147 | -33.824 | 8.1969
0.073251 | -33.8247 | 7.4366
0.079355 | -33.8427 | 6.7714
0.08546 | -33.856 | 6.1748
0.091564 | -33.8508 | 5.7353
0.097668 | -33.8336 | 5.4929
0.10377 | -33.8131 | 5.2531
0.10988 | -33.7969 | 4.9806
0.11598 | -33.7884 | 4.6892
0.12209 | -33.7905 | 4.4033
0.12819 | -33.8005 | 4.1162
0.13429 | -33.8109 | 3.8254
0.1404 | -33.8136 | 3.5457
0.1465 | -33.8154 | 3.2859
0.15261 | -33.8129 | 3.0628
0.15871 | -33.8072 | 2.8969
0.16482 | -33.7997 | 2.7904
0.17092 | -33.7973 | 2.7474
0.17702 | -33.7951 | 2.6721
0.18313 | -33.7941 | 2.5918
0.18923 | -33.7936 | 2.4824
0.19534 | -33.7926 | 2.3564
0.20144 | -33.7918 | 2.2168
0.20754 | -33.79 | 2.0999
0.21365 | -33.7898 | 1.9885
0.21975 | -33.7901 | 1.8935
0.22586 | -33.7917 | 1.8088
0.23196 | -33.7945 | 1.7386
0.23807 | -33.7986 | 1.6778
0.24417 | -33.8021 | 1.6217
0.25027 | -33.8056 | 1.5617
0.25638 | -33.8075 | 1.5025
0.26248 | -33.8087 | 1.4385
0.26859 | -33.8091 | 1.3757
0.27469 | -33.8091 | 1.3071
0.2808 | -33.81 | 1.2357
0.2869 | -33.8107 | 1.1669
0.293 | -33.8115 | 1.1019
0.29911 | -33.8124 | 1.0432
0.30521 | -33.8134 | 0.98527
0.31132 | -33.8142 | 0.94621
0.31742 | -33.8154 | 0.90334
0.32353 | -33.8157 | 0.86787
0.32963 | -33.8166 | 0.83782
0.33573 | -33.8176 | 0.81613
0.34184 | -33.819 | 0.79314
0.34794 | -33.8211 | 0.76789
0.35405 | -33.8232 | 0.74234
0.36015 | -33.8256 | 0.71551
0.36626 | -33.8275 | 0.68677
0.37236 | -33.8295 | 0.64901
0.37846 | -33.8316 | 0.60373
0.38457 | -33.8325 | 0.54569
0.39067 | -33.8336 | 0.49789
0.39678 | -33.8343 | 0.44516
0.40288 | -33.8349 | 0.39026
0.40899 | -33.835 | 0.33755
0.41509 | -33.8346 | 0.28594
0.42119 | -33.8348 | 0.24246
0.4273 | -33.835 | 0.19579
0.4334 | -33.8352 | 0.14962
0.43951 | -33.8357 | 0.10628
0.44561 | -33.8368 | 0.066755
0.45172 | -33.8384 | 0.035802
0.45782 | -33.8393 | -0.00086205
0.46392 | -33.8406 | -0.036455
0.47003 | -33.8418 | -0.068191
0.47613 | -33.8426 | -0.10146
0.48224 | -33.8432 | -0.13525
0.48834 | -33.8421 | -0.16861
0.49445 | -33.8419 | -0.20274
0.50055 | -33.8413 | -0.23349
0.50665 | -33.8405 | -0.2664
0.51276 | -33.8394 | -0.29908
0.51886 | -33.8383 | -0.3317
0.52497 | -33.8371 | -0.35484
0.53107 | -33.8362 | -0.38307
0.53717 | -33.8355 | -0.41008
0.54328 | -33.8346 | -0.43195
0.54938 | -33.8344 | -0.45344
0.55549 | -33.8343 | -0.47367
0.56159 | -33.8354 | -0.48513
0.5677 | -33.8357 | -0.49927
0.5738 | -33.8368 | -0.51528
0.5799 | -33.8376 | -0.53009
0.58601 | -33.8387 | -0.5402
0.59211 | -33.8405 | -0.54682
0.59822 | -33.8418 | -0.55637
0.60432 | -33.8434 | -0.56856
0.61043 | -33.8447 | -0.57669
0.61653 | -33.846 | -0.58703
0.62263 | -33.8476 | -0.5964
0.62874 | -33.8483 | -0.60691
0.63484 | -33.8492 | -0.61631
0.64095 | -33.8497 | -0.6265
0.64705 | -33.8501 | -0.63626
0.65316 | -33.8503 | -0.64858
0.65926 | -33.8504 | -0.66417
0.66536 | -33.8493 | -0.68114
0.67147 | -33.8489 | -0.69405
0.67757 | -33.8486 | -0.708
0.68368 | -33.8479 | -0.72658
0.68978 | -33.8471 | -0.74342
0.69589 | -33.8462 | -0.7621
0.70199 | -33.8453 | -0.7846
0.70809 | -33.8445 | -0.80352
0.7142 | -33.8438 | -0.82382
0.7203 | -33.8427 | -0.84281
0.72641 | -33.8421 | -0.8651
0.73251 | -33.8415 | -0.88499
0.73862 | -33.8418 | -0.90678
0.74472 | -33.8415 | -0.92677
0.75082 | -33.8413 | -0.9488
0.75693 | -33.8416 | -0.97056
0.76303 | -33.8422 | -0.9925
0.76914 | -33.844 | -1.013
0.77524 | -33.845 | -1.0327
0.78135 | -33.8463 | -1.0528
0.78745 | -33.8477 | -1.0727
0.79355 | -33.8489 | -1.0925
0.79966 | -33.8507 | -1.1118
0.80576 | -33.8529 | -1.1313
0.81187 | -33.8545 | -1.1527
0.81797 | -33.8564 | -1.1738
0.82408 | -33.8581 | -1.1942
0.83018 | -33.8598 | -1.2146
0.83628 | -33.8617 | -1.2329
0.84239 | -33.8629 | -1.2532
0.84849 | -33.8643 | -1.2746
0.8546 | -33.8657 | -1.2941
0.8607 | -33.8669 | -1.3148
0.86681 | -33.868 | -1.3335
0.87291 | -33.869 | -1.3521
0.87901 | -33.8694 | -1.3699
0.88512 | -33.8702 | -1.3879
0.89122 | -33.8709 | -1.406
0.89733 | -33.8712 | -1.4215
0.90343 | -33.8717 | -1.4389
0.90953 | -33.8718 | -1.4531
0.91564 | -33.8718 | -1.4688
0.92174 | -33.872 | -1.4836
0.92785 | -33.8721 | -1.498
0.93395 | -33.8722 | -1.5117
0.94006 | -33.8723 | -1.5245
0.94616 | -33.872 | -1.5363
0.95226 | -33.8722 | -1.5488
0.95837 | -33.8722 | -1.5625
0.96447 | -33.8724 | -1.5752
0.97058 | -33.8722 | -1.5885
0.97668 | -33.872 | -1.6017
0.98279 | -33.8717 | -1.6143
0.98889 | -33.8716 | -1.6248
0.99499 | -33.8715 | -1.638
1.0011 | -33.8714 | -1.6485
1.0072 | -33.871 | -1.6601
1.0133 | -33.8707 | -1.6707
1.0194 | -33.8703 | -1.6835
1.0255 | -33.87 | -1.697
1.0316 | -33.8694 | -1.7087
1.0377 | -33.869 | -1.7223
1.0438 | -33.8685 | -1.7348
1.0499 | -33.8677 | -1.75
1.056 | -33.867 | -1.7641
1.0621 | -33.8665 | -1.7782
1.0682 | -33.866 | -1.7906
1.0743 | -33.8654 | -1.8059
1.0805 | -33.8649 | -1.8201
1.0866 | -33.864 | -1.835
1.0927 | -33.8632 | -1.8494
1.0988 | -33.8626 | -1.8642
1.1049 | -33.8618 | -1.8783
1.111 | -33.861 | -1.8934
1.1171 | -33.8603 | -1.9079
1.1232 | -33.8599 | -1.9229
1.1293 | -33.8595 | -1.9368
1.1354 | -33.8592 | -1.9517
1.1415 | -33.8588 | -1.9671
1.1476 | -33.8582 | -1.9824
1.1537 | -33.8577 | -1.9969
1.1598 | -33.8578 | -2.0115
1.1659 | -33.8576 | -2.0247
1.172 | -33.8574 | -2.0385
1.1781 | -33.8571 | -2.0522
1.1842 | -33.857 | -2.0662
1.1903 | -33.8573 | -2.0769
1.1964 | -33.8572 | -2.09
1.2025 | -33.8573 | -2.1037
1.2086 | -33.8577 | -2.116
1.2147 | -33.8577 | -2.1282
1.2209 | -33.858 | -2.1392
1.227 | -33.8587 | -2.1496
1.2331 | -33.8591 | -2.1613
1.2392 | -33.8595 | -2.1728
1.2453 | -33.8599 | -2.1836
1.2514 | -33.8603 | -2.1932
1.2575 | -33.8608 | -2.2039
1.2636 | -33.8616 | -2.2124
1.2697 | -33.8622 | -2.2214
1.2758 | -33.8629 | -2.2321
1.2819 | -33.8636 | -2.2417
1.288 | -33.8642 | -2.2522
1.2941 | -33.865 | -2.2604
1.3002 | -33.8657 | -2.2681
1.3063 | -33.8662 | -2.2769
1.3124 | -33.8669 | -2.2856
1.3185 | -33.8676 | -2.2943
1.3246 | -33.8682 | -2.3031
1.3307 | -33.8691 | -2.3114
1.3368 | -33.8698 | -2.3198
1.3429 | -33.8705 | -2.3284
1.349 | -33.871 | -2.3374
1.3551 | -33.8717 | -2.3464
1.3613 | -33.8722 | -2.3543
1.3674 | -33.8728 | -2.3641
1.3735 | -33.8733 | -2.3736
1.3796 | -33.8737 | -2.3824
1.3857 | -33.874 | -2.3914
1.3918 | -33.8746 | -2.4014
1.3979 | -33.8751 | -2.4103
1.404 | -33.8756 | -2.4215
1.4101 | -33.8761 | -2.4317
1.4162 | -33.8764 | -2.442
1.4223 | -33.8769 | -2.4532
1.4284 | -33.8774 | -2.4631
1.4345 | -33.8777 | -2.4744
1.4406 | -33.878 | -2.4878
1.4467 | -33.8784 | -2.4992
1.4528 | -33.8789 | -2.5104
1.4589 | -33.8794 | -2.5221
1.465 | -33.8796 | -2.5329
1.4711 | -33.8801 | -2.5441
1.4772 | -33.8806 | -2.5573
1.4833 | -33.8811 | -2.5701
1.4894 | -33.8816 | -2.5822
1.4955 | -33.8822 | -2.594
1.5016 | -33.8827 | -2.6064
1.5078 | -33.8831 | -2.6202
1.5139 | -33.8835 | -2.6332
1.52 | -33.884 | -2.6457
1.5261 | -33.8843 | -2.6582
1.5322 | -33.8846 | -2.6721
1.5383 | -33.885 | -2.6846
1.5444 | -33.8856 | -2.6973
1.5505 | -33.886 | -2.7106
1.5566 | -33.8866 | -2.7235
1.5627 | -33.887 | -2.7372
1.5688 | -33.8876 | -2.7507
1.5749 | -33.888 | -2.7633
1.581 | -33.8883 | -2.7766
1.5871 | -33.8887 | -2.7902
1.5932 | -33.889 | -2.803
1.5993 | -33.8894 | -2.8172
1.6054 | -33.8898 | -2.8303
1.6115 | -33.8902 | -2.8444
1.6176 | -33.8905 | -2.8569
1.6237 | -33.8907 | -2.8705
1.6298 | -33.891 | -2.883
1.6359 | -33.8914 | -2.8957
1.642 | -33.8918 | -2.9077
1.6482 | -33.892 | -2.9206
1.6543 | -33.8923 | -2.9342
1.6604 | -33.8926 | -2.9462
1.6665 | -33.8927 | -2.9586
1.6726 | -33.8929 | -2.9711
1.6787 | -33.8932 | -2.9837
1.6848 | -33.8932 | -2.9962
1.6909 | -33.8934 | -3.0089
1.697 | -33.8936 | -3.021
1.7031 | -33.8937 | -3.0336
1.7092 | -33.8936 | -3.0458
1.7153 | -33.8938 | -3.0577
1.7214 | -33.8936 | -3.0698
1.7275 | -33.8935 | -3.0823
1.7336 | -33.8935 | -3.0952
1.7397 | -33.8936 | -3.1075
1.7458 | -33.8936 | -3.1196
1.7519 | -33.8936 | -3.1317
1.758 | -33.8937 | -3.1436
1.7641 | -33.8938 | -3.1551
1.7702 | -33.8938 | -3.168
1.7763 | -33.8937 | -3.1795
1.7824 | -33.8937 | -3.1918
1.7885 | -33.8937 | -3.2041
1.7947 | -33.8938 | -3.216
1.8008 | -33.8938 | -3.2272
1.8069 | -33.8938 | -3.2387
1.813 | -33.8939 | -3.2508
1.8191 | -33.8938 | -3.2634
1.8252 | -33.8939 | -3.276
1.8313 | -33.8941 | -3.2881
1.8374 | -33.8942 | -3.2998
1.8435 | -33.8943 | -3.3117
1.8496 | -33.8943 | -3.3245
1.8557 | -33.8945 | -3.3371
1.8618 | -33.8948 | -3.3491
1.8679 | -33.8949 | -3.361
1.874 | -33.895 | -3.3728
1.8801 | -33.8952 | -3.385
1.8862 | -33.8954 | -3.3968
1.8923 | -33.8955 | -3.4089
1.8984 | -33.8957 | -3.4207
1.9045 | -33.8958 | -3.4321
1.9106 | -33.8958 | -3.4441
1.9167 | -33.8959 | -3.4561
1.9228 | -33.896 | -3.4678
1.9289 | -33.8961 | -3.4795
1.9351 | -33.8963 | -3.4901
1.9412 | -33.8965 | -3.5012
1.9473 | -33.8968 | -3.5131
1.9534 | -33.897 | -3.5244
1.9595 | -33.8972 | -3.536
1.9656 | -33.8975 | -3.548
1.9717 | -33.8977 | -3.5595
1.9778 | -33.8981 | -3.5707
1.9839 | -33.8983 | -3.582
1.99 | -33.8985 | -3.5928
1.9961 | -33.8986 | -3.6048
2.0022 | -33.8988 | -3.616
2.0083 | -33.8989 | -3.6271
2.0144 | -33.8991 | -3.6383
2.0205 | -33.8995 | -3.6499
2.0266 | -33.8998 | -3.6611
2.0327 | -33.9 | -3.6724
2.0388 | -33.9004 | -3.6838
2.0449 | -33.9006 | -3.695
2.051 | -33.9009 | -3.7062
2.0571 | -33.9012 | -3.7174
2.0632 | -33.9014 | -3.7287
2.0693 | -33.9016 | -3.7406
2.0754 | -33.9019 | -3.7519
2.0816 | -33.9022 | -3.7638
2.0877 | -33.9024 | -3.7751
2.0938 | -33.9027 | -3.7871
2.0999 | -33.9029 | -3.7986
2.106 | -33.9031 | -3.8101
2.1121 | -33.9033 | -3.8219
2.1182 | -33.9035 | -3.8326
2.1243 | -33.9038 | -3.8432
2.1304 | -33.9039 | -3.8542
2.1365 | -33.9042 | -3.8658
2.1426 | -33.9044 | -3.8776
2.1487 | -33.9047 | -3.8887
2.1548 | -33.9051 | -3.9004
2.1609 | -33.9055 | -3.9117
2.167 | -33.9059 | -3.9227
2.1731 | -33.9062 | -3.9346
2.1792 | -33.9064 | -3.9455
2.1853 | -33.9068 | -3.9573
2.1914 | -33.907 | -3.9685
2.1975 | -33.9074 | -3.9798
2.2036 | -33.9076 | -3.9917
2.2097 | -33.9078 | -4.003
2.2158 | -33.9082 | -4.0135
2.222 | -33.9084 | -4.0239
2.2281 | -33.9087 | -4.0351
2.2342 | -33.9089 | -4.0463
2.2403 | -33.9091 | -4.0568
2.2464 | -33.9094 | -4.0675
2.2525 | -33.9096 | -4.0788
2.2586 | -33.9099 | -4.0892
2.2647 | -33.9102 | -4.0995
2.2708 | -33.9105 | -4.1106
2.2769 | -33.9108 | -4.1212
2.283 | -33.911 | -4.1322
2.2891 | -33.9112 | -4.1429
2.2952 | -33.9113 | -4.1534
2.3013 | -33.9115 | -4.1641
2.3074 | -33.9118 | -4.1744
2.3135 | -33.9119 | -4.1857
2.3196 | -33.9121 | -4.1957
2.3257 | -33.9123 | -4.2062
2.3318 | -33.9124 | -4.2171
2.3379 | -33.9126 | -4.2277
2.344 | -33.9127 | -4.239
2.3501 | -33.9129 | -4.2495
2.3562 | -33.9131 | -4.2598
2.3623 | -33.9132 | -4.2701
2.3685 | -33.9133 | -4.2805
2.3746 | -33.9134 | -4.2916
2.3807 | -33.9136 | -4.3019
2.3868 | -33.9137 | -4.3125
2.3929 | -33.9139 | -4.323
2.399 | -33.914 | -4.3333
2.4051 | -33.9142 | -4.3437
2.4112 | -33.9144 | -4.3543
2.4173 | -33.9145 | -4.3647
2.4234 | -33.9146 | -4.3754
2.4295 | -33.9147 | -4.3861
2.4356 | -33.9149 | -4.3968
2.4417 | -33.915 | -4.4078
2.4478 | -33.9151 | -4.4185
2.4539 | -33.9152 | -4.4286
2.46 | -33.9153 | -4.4399
2.4661 | -33.9155 | -4.4509
2.4722 | -33.9156 | -4.461
2.4783 | -33.9157 | -4.4713
2.4844 | -33.9158 | -4.4822
2.4905 | -33.9158 | -4.493
2.4966 | -33.916 | -4.5036
2.5027 | -33.9161 | -4.5139
2.5089 | -33.9162 | -4.5247
2.515 | -33.9162 | -4.5353
2.5211 | -33.9163 | -4.5459
2.5272 | -33.9164 | -4.5568
2.5333 | -33.9165 | -4.5676
2.5394 | -33.9167 | -4.5782
2.5455 | -33.9168 | -4.5894
2.5516 | -33.9169 | -4.5996
2.5577 | -33.9171 | -4.611
2.5638 | -33.9173 | -4.6226
2.5699 | -33.9173 | -4.6339
2.576 | -33.9173 | -4.6453
2.5821 | -33.9175 | -4.6565
2.5882 | -33.9176 | -4.6675
2.5943 | -33.9177 | -4.6782
2.6004 | -33.9179 | -4.689
2.6065 | -33.918 | -4.6993
2.6126 | -33.9181 | -4.7097
2.6187 | -33.9182 | -4.7201
2.6248 | -33.9184 | -4.7308
2.6309 | -33.9185 | -4.7421
2.637 | -33.9188 | -4.7536
2.6431 | -33.9189 | -4.764
2.6492 | -33.919 | -4.7753
2.6554 | -33.9192 | -4.786
2.6615 | -33.9193 | -4.7969
2.6676 | -33.9196 | -4.8082
2.6737 | -33.9199 | -4.8193
2.6798 | -33.9202 | -4.8303
2.6859 | -33.9203 | -4.8415
2.692 | -33.9205 | -4.8523
2.6981 | -33.9207 | -4.8633
2.7042 | -33.9209 | -4.8744
2.7103 | -33.9211 | -4.8849
2.7164 | -33.9214 | -4.8958
2.7225 | -33.9217 | -4.907
2.7286 | -33.9219 | -4.918
2.7347 | -33.9221 | -4.9284
2.7408 | -33.9224 | -4.9384
2.7469 | -33.9226 | -4.9496
2.753 | -33.9229 | -4.9601
2.7591 | -33.9232 | -4.9713
2.7652 | -33.9234 | -4.9827
2.7713 | -33.9236 | -4.9934
2.7774 | -33.9239 | -5.0042
2.7835 | -33.9241 | -5.0151
2.7896 | -33.9243 | -5.0255
2.7958 | -33.9245 | -5.0362
2.8019 | -33.9249 | -5.0465
2.808 | -33.9251 | -5.0573
2.8141 | -33.9254 | -5.0684
2.8202 | -33.9257 | -5.0791
2.8263 | -33.926 | -5.09
2.8324 | -33.9264 | -5.1012
2.8385 | -33.9268 | -5.112
2.8446 | -33.9272 | -5.1227
2.8507 | -33.9275 | -5.1334
2.8568 | -33.9279 | -5.1438
2.8629 | -33.9281 | -5.1543
2.869 | -33.9285 | -5.1653
2.8751 | -33.9288 | -5.176
2.8812 | -33.9292 | -5.1872
2.8873 | -33.9295 | -5.198
2.8934 | -33.9298 | -5.2089
2.8995 | -33.9302 | -5.2199
2.9056 | -33.9304 | -5.2309
2.9117 | -33.9308 | -5.2418
2.9178 | -33.9311 | -5.2525
2.9239 | -33.9315 | -5.2639
2.93 | -33.9319 | -5.2752
2.9361 | -33.9322 | -5.2859
2.9423 | -33.9326 | -5.2965
2.9484 | -33.933 | -5.3075
2.9545 | -33.9333 | -5.3184
2.9606 | -33.9337 | -5.3296
2.9667 | -33.9341 | -5.3406
2.9728 | -33.9344 | -5.3518
2.9789 | -33.9349 | -5.3628
2.985 | -33.9353 | -5.3739
2.9911 | -33.9356 | -5.385
2.9972 | -33.9361 | -5.396
3.0033 | -33.9365 | -5.4074
3.0094 | -33.9368 | -5.4183
3.0155 | -33.9373 | -5.4291
3.0216 | -33.9376 | -5.4402
3.0277 | -33.9379 | -5.4517
3.0338 | -33.9383 | -5.463
3.0399 | -33.9387 | -5.4741
3.046 | -33.939 | -5.4855
3.0521 | -33.9394 | -5.4965
3.0582 | -33.9396 | -5.5083
3.0643 | -33.94 | -5.5196
3.0704 | -33.9403 | -5.5312
3.0765 | -33.9406 | -5.5422
3.0827 | -33.9409 | -5.5532
3.0888 | -33.9412 | -5.5643
3.0949 | -33.9416 | -5.5755
3.101 | -33.9419 | -5.5868
3.1071 | -33.9422 | -5.5983
3.1132 | -33.9426 | -5.6099
3.1193 | -33.9429 | -5.6212
3.1254 | -33.9433 | -5.6324
3.1315 | -33.9436 | -5.6441
3.1376 | -33.9439 | -5.6553
3.1437 | -33.9442 | -5.6668
3.1498 | -33.9446 | -5.6786
3.1559 | -33.9449 | -5.6899
3.162 | -33.9451 | -5.7017
3.1681 | -33.9454 | -5.7135
3.1742 | -33.9457 | -5.7248
3.1803 | -33.946 | -5.7363
3.1864 | -33.9463 | -5.7474
3.1925 | -33.9466 | -5.7587
3.1986 | -33.9468 | -5.7705
3.2047 | -33.9471 | -5.7823
3.2108 | -33.9473 | -5.7932
3.2169 | -33.9476 | -5.8047
3.223 | -33.9478 | -5.8162
3.2292 | -33.9481 | -5.8279
3.2353 | -33.9483 | -5.8385
3.2414 | -33.9485 | -5.8497
3.2475 | -33.9487 | -5.861
3.2536 | -33.949 | -5.8722
3.2597 | -33.9492 | -5.883
3.2658 | -33.9494 | -5.8938
3.2719 | -33.9496 | -5.9047
3.278 | -33.9498 | -5.9161
3.2841 | -33.9501 | -5.927
3.2902 | -33.9503 | -5.9382
3.2963 | -33.9505 | -5.9495
3.3024 | -33.9506 | -5.9604
3.3085 | -33.9509 | -5.9718
3.3146 | -33.951 | -5.9828
3.3207 | -33.9513 | -5.9942
3.3268 | -33.9515 | -6.0054
3.3329 | -33.9517 | -6.016
3.339 | -33.9518 | -6.0266
3.3451 | -33.952 | -6.0376
3.3512 | -33.9521 | -6.0482
3.3573 | -33.9524 | -6.0592
3.3634 | -33.9525 | -6.0702
3.3696 | -33.9527 | -6.0814
3.3757 | -33.9528 | -6.092
3.3818 | -33.9531 | -6.1027
3.3879 | -33.9532 | -6.1137
3.394 | -33.9534 | -6.1249
3.4001 | -33.9536 | -6.1356
3.4062 | -33.9536 | -6.1464
3.4123 | -33.9538 | -6.1568
3.4184 | -33.954 | -6.1679
3.4245 | -33.9542 | -6.1786
3.4306 | -33.9543 | -6.1891
3.4367 | -33.9545 | -6.1996
3.4428 | -33.9546 | -6.2098
3.4489 | -33.9547 | -6.2206
3.455 | -33.9549 | -6.2314
3.4611 | -33.9551 | -6.2425
3.4672 | -33.9552 | -6.2534
3.4733 | -33.9554 | -6.2643
3.4794 | -33.9555 | -6.2751
3.4855 | -33.9557 | -6.2854
3.4916 | -33.9559 | -6.2957
3.4977 | -33.956 | -6.3068
3.5038 | -33.9561 | -6.3176
3.5099 | -33.9563 | -6.3277
3.5161 | -33.9564 | -6.338
3.5222 | -33.9565 | -6.3483
3.5283 | -33.9566 | -6.359
3.5344 | -33.9568 | -6.3697
3.5405 | -33.9569 | -6.3802
3.5466 | -33.9571 | -6.391
3.5527 | -33.9572 | -6.4018
3.5588 | -33.9573 | -6.4123
3.5649 | -33.9574 | -6.4229
3.571 | -33.9575 | -6.4334
3.5771 | -33.9576 | -6.4444
3.5832 | -33.9578 | -6.4542
3.5893 | -33.958 | -6.4648
3.5954 | -33.9581 | -6.4755
3.6015 | -33.9583 | -6.4863
3.6076 | -33.9585 | -6.4967
3.6137 | -33.9587 | -6.5068
3.6198 | -33.9589 | -6.517
3.6259 | -33.9591 | -6.5278
3.632 | -33.9592 | -6.5384
3.6381 | -33.9594 | -6.5491
3.6442 | -33.9596 | -6.5596
3.6503 | -33.9598 | -6.5697
3.6565 | -33.96 | -6.5804
3.6626 | -33.9601 | -6.5908
3.6687 | -33.9604 | -6.6015
3.6748 | -33.9606 | -6.6118
3.6809 | -33.9609 | -6.6223
3.687 | -33.961 | -6.6332
3.6931 | -33.9613 | -6.6434
3.6992 | -33.9615 | -6.654
3.7053 | -33.9617 | -6.6645
3.7114 | -33.9618 | -6.6751
3.7175 | -33.962 | -6.6857
3.7236 | -33.9623 | -6.6957
3.7297 | -33.9625 | -6.7065
3.7358 | -33.9627 | -6.7173
3.7419 | -33.9629 | -6.728
3.748 | -33.9631 | -6.7385
3.7541 | -33.9633 | -6.7491
3.7602 | -33.9636 | -6.7595
3.7663 | -33.9638 | -6.7699
3.7724 | -33.9641 | -6.7806
3.7785 | -33.9643 | -6.7909
3.7846 | -33.9645 | -6.8013
3.7907 | -33.9648 | -6.8122
3.7969 | -33.9651 | -6.8229
3.803 | -33.9653 | -6.833
3.8091 | -33.9655 | -6.8435
3.8152 | -33.9658 | -6.854
3.8213 | -33.966 | -6.8648
3.8274 | -33.9663 | -6.8751
3.8335 | -33.9666 | -6.8858
3.8396 | -33.9669 | -6.8964
3.8457 | -33.9672 | -6.907
3.8518 | -33.9675 | -6.9175
3.8579 | -33.9679 | -6.928
3.864 | -33.9681 | -6.9387
3.8701 | -33.9684 | -6.9496
3.8762 | -33.9687 | -6.96
3.8823 | -33.9689 | -6.9708
3.8884 | -33.9692 | -6.9814
3.8945 | -33.9696 | -6.9922
3.9006 | -33.9698 | -7.0031
3.9067 | -33.9701 | -7.0142
3.9128 | -33.9704 | -7.0248
3.9189 | -33.9707 | -7.0353
3.925 | -33.9709 | -7.046
3.9311 | -33.9712 | -7.0562
3.9372 | -33.9715 | -7.067
3.9434 | -33.9718 | -7.0776
3.9495 | -33.9721 | -7.0883
3.9556 | -33.9724 | -7.0988
3.9617 | -33.9727 | -7.109
3.9678 | -33.973 | -7.1193
3.9739 | -33.9733 | -7.1299
3.98 | -33.9735 | -7.1406
3.9861 | -33.9738 | -7.1512
3.9922 | -33.974 | -7.1618
3.9983 | -33.9743 | -7.172
4.0044 | -33.9746 | -7.1822
4.0105 | -33.9749 | -7.1928
4.0166 | -33.9752 | -7.2035
4.0227 | -33.9755 | -7.214
4.0288 | -33.9757 | -7.225
4.0349 | -33.9759 | -7.2354
4.041 | -33.9762 | -7.2459
4.0471 | -33.9764 | -7.2566
4.0532 | -33.9767 | -7.2668
4.0593 | -33.977 | -7.2776
4.0654 | -33.9773 | -7.2879
4.0715 | -33.9775 | -7.2989
4.0776 | -33.9778 | -7.3094
4.0838 | -33.9781 | -7.3205
4.0899 | -33.9784 | -7.3308
4.096 | -33.9787 | -7.3413
4.1021 | -33.9789 | -7.3516
4.1082 | -33.9792 | -7.3619
4.1143 | -33.9794 | -7.3721
4.1204 | -33.9797 | -7.3824
4.1265 | -33.9799 | -7.3926
4.1326 | -33.9802 | -7.4029
4.1387 | -33.9805 | -7.4133
4.1448 | -33.9807 | -7.4239
4.1509 | -33.9809 | -7.4344
4.157 | -33.9812 | -7.4453
4.1631 | -33.9815 | -7.456
4.1692 | -33.9817 | -7.4664
4.1753 | -33.982 | -7.4768
4.1814 | -33.9822 | -7.487
4.1875 | -33.9825 | -7.4977
4.1936 | -33.9827 | -7.5082
4.1997 | -33.9829 | -7.5183
4.2058 | -33.9832 | -7.5291
4.2119 | -33.9834 | -7.5394
4.218 | -33.9836 | -7.5496
4.2241 | -33.9838 | -7.5598
4.2303 | -33.9841 | -7.5701
4.2364 | -33.9843 | -7.5805
4.2425 | -33.9845 | -7.591
4.2486 | -33.9847 | -7.6015
4.2547 | -33.9848 | -7.6117
4.2608 | -33.985 | -7.6222
4.2669 | -33.9852 | -7.6323
4.273 | -33.9855 | -7.6428
4.2791 | -33.9857 | -7.6532
4.2852 | -33.9859 | -7.6638
4.2913 | -33.986 | -7.6739
4.2974 | -33.9863 | -7.6846
4.3035 | -33.9865 | -7.6953
4.3096 | -33.9867 | -7.7058
4.3157 | -33.9869 | -7.716
4.3218 | -33.9871 | -7.7263
4.3279 | -33.9873 | -7.7366
4.334 | -33.9875 | -7.7469
4.3401 | -33.9877 | -7.7569
4.3462 | -33.9879 | -7.7674
4.3523 | -33.9882 | -7.7773
4.3584 | -33.9883 | -7.7872
4.3645 | -33.9885 | -7.7975
4.3707 | -33.9887 | -7.8075
4.3768 | -33.9889 | -7.8178
4.3829 | -33.9892 | -7.8283
4.389 | -33.9894 | -7.8386
4.3951 | -33.9896 | -7.8487
4.4012 | -33.9897 | -7.8592
4.4073 | -33.9899 | -7.8693
4.4134 | -33.9901 | -7.8795
4.4195 | -33.9904 | -7.8898
4.4256 | -33.9906 | -7.8998
4.4317 | -33.9908 | -7.91
4.4378 | -33.991 | -7.9199
4.4439 | -33.9912 | -7.9298
4.45 | -33.9913 | -7.94
4.4561 | -33.9915 | -7.9502
4.4622 | -33.9917 | -7.9604
4.4683 | -33.9919 | -7.9706
4.4744 | -33.9921 | -7.9812
4.4805 | -33.9923 | -7.9909
4.4866 | -33.9925 | -8.0011
4.4927 | -33.9927 | -8.0115
4.4988 | -33.9929 | -8.0217
4.5049 | -33.9931 | -8.0315
4.511 | -33.9933 | -8.0414
4.5172 | -33.9935 | -8.0515
4.5233 | -33.9937 | -8.0621
4.5294 | -33.994 | -8.0721
4.5355 | -33.9942 | -8.0822
4.5416 | -33.9944 | -8.0924
4.5477 | -33.9946 | -8.1025
4.5538 | -33.9948 | -8.1123
4.5599 | -33.995 | -8.1224
4.566 | -33.9952 | -8.1325
4.5721 | -33.9954 | -8.1427
4.5782 | -33.9957 | -8.1531
4.5843 | -33.9959 | -8.1633
4.5904 | -33.9961 | -8.1735
4.5965 | -33.9964 | -8.1835
4.6026 | -33.9966 | -8.194
4.6087 | -33.9968 | -8.2041
4.6148 | -33.997 | -8.2145
4.6209 | -33.9972 | -8.2246
4.627 | -33.9974 | -8.2349
4.6331 | -33.9976 | -8.245
4.6392 | -33.9978 | -8.2552
4.6453 | -33.998 | -8.2654
4.6514 | -33.9983 | -8.2755
4.6576 | -33.9985 | -8.2861
4.6637 | -33.9987 | -8.2959
4.6698 | -33.9988 | -8.3059
4.6759 | -33.9991 | -8.3164
4.682 | -33.9993 | -8.3265
4.6881 | -33.9994 | -8.3367
4.6942 | -33.9997 | -8.3468
4.7003 | -33.9999 | -8.357
4.7064 | -34.0001 | -8.367
4.7125 | -34.0003 | -8.3775
4.7186 | -34.0005 | -8.3875
4.7247 | -34.0008 | -8.3977
4.7308 | -34.001 | -8.4078
4.7369 | -34.0012 | -8.4181
4.743 | -34.0015 | -8.4287
4.7491 | -34.0017 | -8.4392
4.7552 | -34.0019 | -8.4495
4.7613 | -34.0021 | -8.4598
4.7674 | -34.0023 | -8.4699
4.7735 | -34.0025 | -8.4805
4.7796 | -34.0028 | -8.491
4.7857 | -34.003 | -8.5013
4.7918 | -34.0032 | -8.5113
4.7979 | -34.0035 | -8.5214
4.8041 | -34.0037 | -8.5317
4.8102 | -34.004 | -8.5416
4.8163 | -34.0042 | -8.5518
4.8224 | -34.0044 | -8.562
4.8285 | -34.0047 | -8.572
4.8346 | -34.0049 | -8.582
4.8407 | -34.0051 | -8.5923
4.8468 | -34.0054 | -8.6028
4.8529 | -34.0056 | -8.6133
4.859 | -34.0058 | -8.6239
4.8651 | -34.0061 | -8.6339
4.8712 | -34.0063 | -8.644
4.8773 | -34.0065 | -8.6543
4.8834 | -34.0067 | -8.6645
4.8895 | -34.007 | -8.6746
4.8956 | -34.0072 | -8.6849
4.9017 | -34.0074 | -8.695
4.9078 | -34.0076 | -8.7055
4.9139 | -34.0079 | -8.7154
4.92 | -34.0082 | -8.7254
4.9261 | -34.0084 | -8.7353
4.9322 | -34.0087 | -8.7455
4.9383 | -34.0089 | -8.7557
4.9445 | -34.0092 | -8.7659
4.9506 | -34.0094 | -8.776
4.9567 | -34.0096 | -8.7861
4.9628 | -34.0099 | -8.7963
4.9689 | -34.0101 | -8.8069
4.975 | -34.0103 | -8.8174
4.9811 | -34.0106 | -8.8273
4.9872 | -34.0109 | -8.8375
4.9933 | -34.0112 | -8.848
4.9994 | -34.0115 | -8.8586
5.0055 | -34.0117 | -8.8689
5.0116 | -34.012 | -8.8795
5.0177 | -34.0123 | -8.8901
5.0238 | -34.0126 | -8.9006
5.0299 | -34.0129 | -8.9112
5.036 | -34.0132 | -8.9215
5.0421 | -34.0134 | -8.9316
5.0482 | -34.0137 | -8.9419
5.0543 | -34.014 | -8.952
5.0604 | -34.0143 | -8.9623
5.0665 | -34.0146 | -8.9723
5.0726 | -34.0148 | -8.9829
5.0787 | -34.0151 | -8.9934
5.0848 | -34.0154 | -9.0038
5.091 | -34.0157 | -9.0139
5.0971 | -34.016 | -9.0239
5.1032 | -34.0162 | -9.0345
5.1093 | -34.0165 | -9.0448
5.1154 | -34.0168 | -9.0549
5.1215 | -34.017 | -9.065
5.1276 | -34.0173 | -9.0751
5.1337 | -34.0176 | -9.0852
5.1398 | -34.0179 | -9.0952
5.1459 | -34.0182 | -9.1058
5.152 | -34.0185 | -9.1162
5.1581 | -34.0187 | -9.1267
5.1642 | -34.019 | -9.137
5.1703 | -34.0193 | -9.1473
5.1764 | -34.0196 | -9.1574
5.1825 | -34.0199 | -9.1674
5.1886 | -34.0202 | -9.1774
5.1947 | -34.0204 | -9.1877
5.2008 | -34.0207 | -9.1981
5.2069 | -34.021 | -9.2085
5.213 | -34.0213 | -9.2186
5.2191 | -34.0216 | -9.229
5.2252 | -34.0219 | -9.2393
5.2314 | -34.0221 | -9.2498
5.2375 | -34.0224 | -9.2598
5.2436 | -34.0227 | -9.2701
5.2497 | -34.023 | -9.2801
5.2558 | -34.0233 | -9.2908
5.2619 | -34.0236 | -9.3013
5.268 | -34.0239 | -9.3118
5.2741 | -34.0242 | -9.3221
5.2802 | -34.0245 | -9.3328
5.2863 | -34.0248 | -9.343
5.2924 | -34.0251 | -9.3535
5.2985 | -34.0255 | -9.3639
5.3046 | -34.0258 | -9.3742
5.3107 | -34.026 | -9.3847
5.3168 | -34.0263 | -9.3951
5.3229 | -34.0266 | -9.4055
5.329 | -34.0269 | -9.4158
5.3351 | -34.0272 | -9.4258
5.3412 | -34.0275 | -9.4361
5.3473 | -34.0278 | -9.4465
5.3534 | -34.0281 | -9.4569
5.3595 | -34.0284 | -9.4673
5.3656 | -34.0287 | -9.4776
5.3717 | -34.029 | -9.4882
5.3779 | -34.0293 | -9.4989
5.384 | -34.0296 | -9.5092
5.3901 | -34.0299 | -9.5198
5.3962 | -34.0302 | -9.5299
5.4023 | -34.0305 | -9.5405
5.4084 | -34.0309 | -9.5507
5.4145 | -34.0312 | -9.5608
5.4206 | -34.0314 | -9.5712
5.4267 | -34.0318 | -9.5819
5.4328 | -34.0321 | -9.5923
5.4389 | -34.0324 | -9.6023
5.445 | -34.0326 | -9.613
5.4511 | -34.0329 | -9.6232
5.4572 | -34.0333 | -9.6336
5.4633 | -34.0336 | -9.644
5.4694 | -34.0338 | -9.6544
5.4755 | -34.0341 | -9.6645
5.4816 | -34.0344 | -9.6749
5.4877 | -34.0347 | -9.6851
5.4938 | -34.035 | -9.6952
5.4999 | -34.0353 | -9.7054
5.506 | -34.0357 | -9.7157
5.5121 | -34.036 | -9.726
5.5183 | -34.0363 | -9.736
5.5244 | -34.0365 | -9.7466
5.5305 | -34.0369 | -9.7568
5.5366 | -34.0371 | -9.7669
5.5427 | -34.0374 | -9.7774
5.5488 | -34.0377 | -9.7877
5.5549 | -34.038 | -9.798
5.561 | -34.0383 | -9.8086
5.5671 | -34.0386 | -9.8187
5.5732 | -34.0389 | -9.8293
5.5793 | -34.0392 | -9.8399
5.5854 | -34.0396 | -9.8501
5.5915 | -34.0398 | -9.8608
5.5976 | -34.0401 | -9.8712
5.6037 | -34.0404 | -9.8815
5.6098 | -34.0407 | -9.8918
5.6159 | -34.041 | -9.902
5.622 | -34.0413 | -9.9124
5.6281 | -34.0416 | -9.9227
5.6342 | -34.0419 | -9.9334
5.6403 | -34.0422 | -9.9438
5.6464 | -34.0425 | -9.9542
5.6525 | -34.0428 | -9.9647
5.6586 | -34.0431 | -9.9747
5.6648 | -34.0434 | -9.9851
5.6709 | -34.0437 | -9.9956
5.677 | -34.044 | -10.006
5.6831 | -34.0443 | -10.0164
5.6892 | -34.0446 | -10.0269
5.6953 | -34.0449 | -10.0372
5.7014 | -34.0452 | -10.0476
5.7075 | -34.0455 | -10.0579
5.7136 | -34.0457 | -10.0682
5.7197 | -34.046 | -10.0785
5.7258 | -34.0463 | -10.0886
5.7319 | -34.0466 | -10.099
5.738 | -34.0469 | -10.1093
5.7441 | -34.0472 | -10.1196
5.7502 | -34.0476 | -10.1297
5.7563 | -34.0479 | -10.14
5.7624 | -34.0482 | -10.1505
5.7685 | -34.0485 | -10.1609
5.7746 | -34.0487 | -10.1715
5.7807 | -34.049 | -10.1819
5.7868 | -34.0493 | -10.1924
5.7929 | -34.0495 | -10.2026
5.799 | -34.0498 | -10.2128
5.8052 | -34.0501 | -10.2229
5.8113 | -34.0504 | -10.2331
5.8174 | -34.0507 | -10.2433
5.8235 | -34.051 | -10.2536
5.8296 | -34.0513 | -10.2641
5.8357 | -34.0515 | -10.2743
5.8418 | -34.0518 | -10.2847
5.8479 | -34.0521 | -10.2951
5.854 | -34.0524 | -10.3056
5.8601 | -34.0527 | -10.3158
5.8662 | -34.053 | -10.3263
5.8723 | -34.0532 | -10.3364
5.8784 | -34.0535 | -10.3466
5.8845 | -34.0538 | -10.3568
5.8906 | -34.0541 | -10.3672
5.8967 | -34.0544 | -10.3773
5.9028 | -34.0547 | -10.3874
5.9089 | -34.0551 | -10.3976
5.915 | -34.0553 | -10.4076
5.9211 | -34.0556 | -10.4178
5.9272 | -34.0559 | -10.4279
5.9333 | -34.0562 | -10.438
5.9394 | -34.0565 | -10.4482
5.9455 | -34.0568 | -10.4583
5.9517 | -34.0572 | -10.4684
5.9578 | -34.0575 | -10.4788
5.9639 | -34.0578 | -10.4887
5.97 | -34.0581 | -10.4991
5.9761 | -34.0584 | -10.5091
5.9822 | -34.0587 | -10.5195
5.9883 | -34.059 | -10.5298
5.9944 | -34.0593 | -10.5401
6.0005 | -34.0596 | -10.5505
6.0066 | -34.0599 | -10.5606
6.0127 | -34.0603 | -10.5705
6.0188 | -34.0606 | -10.581
6.0249 | -34.0608 | -10.5913
6.031 | -34.0611 | -10.6013
6.0371 | -34.0614 | -10.6115
6.0432 | -34.0617 | -10.6217
6.0493 | -34.062 | -10.6317
6.0554 | -34.0624 | -10.6418
6.0615 | -34.0627 | -10.652
6.0676 | -34.063 | -10.6624
6.0737 | -34.0633 | -10.6727
6.0798 | -34.0636 | -10.6829
6.0859 | -34.0639 | -10.6931
6.0921 | -34.0642 | -10.703
6.0982 | -34.0645 | -10.7132
6.1043 | -34.0648 | -10.7233
6.1104 | -34.0652 | -10.7337
6.1165 | -34.0655 | -10.7435
6.1226 | -34.0658 | -10.7536
6.1287 | -34.0661 | -10.7637
6.1348 | -34.0664 | -10.7738
6.1409 | -34.0667 | -10.7839
6.147 | -34.067 | -10.7941
6.1531 | -34.0673 | -10.8042
6.1592 | -34.0676 | -10.8142
6.1653 | -34.0679 | -10.8243
6.1714 | -34.0682 | -10.8345
6.1775 | -34.0685 | -10.8448
6.1836 | -34.0688 | -10.8549
6.1897 | -34.0692 | -10.8651
6.1958 | -34.0695 | -10.8755
6.2019 | -34.0698 | -10.8857
6.208 | -34.0701 | -10.8959
6.2141 | -34.0704 | -10.906
6.2202 | -34.0708 | -10.9162
6.2263 | -34.0711 | -10.9263
6.2325 | -34.0714 | -10.9362
6.2386 | -34.0717 | -10.9461
6.2447 | -34.072 | -10.9565
6.2508 | -34.0723 | -10.9666
6.2569 | -34.0727 | -10.9767
6.263 | -34.073 | -10.9868
6.2691 | -34.0733 | -10.9968
6.2752 | -34.0735 | -11.007
6.2813 | -34.0738 | -11.0171
6.2874 | -34.0742 | -11.0271
6.2935 | -34.0745 | -11.0372
6.2996 | -34.0747 | -11.0473
6.3057 | -34.0751 | -11.0575
6.3118 | -34.0753 | -11.0677
6.3179 | -34.0756 | -11.0778
6.324 | -34.076 | -11.0879
6.3301 | -34.0763 | -11.0981
6.3362 | -34.0766 | -11.1081
6.3423 | -34.0769 | -11.1184
6.3484 | -34.0772 | -11.1286
6.3545 | -34.0775 | -11.1384
6.3606 | -34.0778 | -11.1483
6.3667 | -34.0781 | -11.1583
6.3728 | -34.0784 | -11.1684
6.379 | -34.0787 | -11.1785
6.3851 | -34.0791 | -11.1884
6.3912 | -34.0793 | -11.1984
6.3973 | -34.0796 | -11.2086
6.4034 | -34.0799 | -11.2185
6.4095 | -34.0802 | -11.2286
6.4156 | -34.0805 | -11.2387
6.4217 | -34.0808 | -11.2487
6.4278 | -34.0812 | -11.2589
6.4339 | -34.0815 | -11.2689
6.44 | -34.0818 | -11.2788
6.4461 | -34.0821 | -11.2891
6.4522 | -34.0824 | -11.299
6.4583 | -34.0827 | -11.3091
6.4644 | -34.083 | -11.3193
6.4705 | -34.0833 | -11.3291
6.4766 | -34.0836 | -11.3392
6.4827 | -34.0839 | -11.3493
6.4888 | -34.0842 | -11.3592
6.4949 | -34.0845 | -11.369
6.501 | -34.0848 | -11.3792
6.5071 | -34.0851 | -11.3891
6.5132 | -34.0854 | -11.3992
6.5194 | -34.0857 | -11.4093
6.5255 | -34.0861 | -11.4191
6.5316 | -34.0863 | -11.4293
6.5377 | -34.0866 | -11.4393
6.5438 | -34.0869 | -11.4492
6.5499 | -34.0872 | -11.4592
6.556 | -34.0876 | -11.4692
6.5621 | -34.0878 | -11.4794
6.5682 | -34.0882 | -11.4895
6.5743 | -34.0885 | -11.4995
6.5804 | -34.0888 | -11.5097
6.5865 | -34.0891 | -11.5198
6.5926 | -34.0894 | -11.5297
6.5987 | -34.0897 | -11.5395
6.6048 | -34.09 | -11.5496
6.6109 | -34.0904 | -11.5597
6.617 | -34.0906 | -11.5699
6.6231 | -34.0909 | -11.5799
6.6292 | -34.0913 | -11.59
6.6353 | -34.0916 | -11.6003
6.6414 | -34.0919 | -11.6102
6.6475 | -34.0922 | -11.6203
6.6536 | -34.0925 | -11.6303
6.6597 | -34.0928 | -11.6401
6.6659 | -34.0931 | -11.6501
6.672 | -34.0934 | -11.6601
6.6781 | -34.0937 | -11.6702
6.6842 | -34.0941 | -11.68
6.6903 | -34.0944 | -11.69
6.6964 | -34.0947 | -11.7
6.7025 | -34.095 | -11.7098
6.7086 | -34.0953 | -11.7198
6.7147 | -34.0956 | -11.7299
6.7208 | -34.0959 | -11.7401
6.7269 | -34.0962 | -11.75
6.733 | -34.0965 | -11.7598
6.7391 | -34.0968 | -11.7699
6.7452 | -34.0972 | -11.7797
6.7513 | -34.0975 | -11.7895
6.7574 | -34.0978 | -11.7995
6.7635 | -34.0981 | -11.8093
6.7696 | -34.0984 | -11.8192
6.7757 | -34.0988 | -11.8292
6.7818 | -34.0991 | -11.8393
6.7879 | -34.0994 | -11.8491
6.794 | -34.0997 | -11.859
6.8001 | -34.1 | -11.8689
6.8063 | -34.1003 | -11.8789
6.8124 | -34.1006 | -11.8891
6.8185 | -34.1009 | -11.8992
6.8246 | -34.1012 | -11.9092
6.8307 | -34.1015 | -11.919
6.8368 | -34.1018 | -11.9292
6.8429 | -34.1021 | -11.9393
6.849 | -34.1024 | -11.9493
6.8551 | -34.1028 | -11.9592
6.8612 | -34.1031 | -11.9691
6.8673 | -34.1034 | -11.9791
6.8734 | -34.1037 | -11.9892
6.8795 | -34.104 | -11.9992
6.8856 | -34.1043 | -12.0092
6.8917 | -34.1047 | -12.0191
6.8978 | -34.1049 | -12.0289
6.9039 | -34.1053 | -12.0386
6.91 | -34.1056 | -12.0486
6.9161 | -34.1059 | -12.0584
6.9222 | -34.1062 | -12.0684
6.9283 | -34.1066 | -12.0781
6.9344 | -34.1069 | -12.088
6.9405 | -34.1072 | -12.0979
6.9466 | -34.1076 | -12.108
6.9528 | -34.1079 | -12.1182
6.9589 | -34.1083 | -12.1283
6.965 | -34.1086 | -12.1383
6.9711 | -34.1089 | -12.1484
6.9772 | -34.1093 | -12.1583
6.9833 | -34.1096 | -12.1682
6.9894 | -34.1099 | -12.1784
6.9955 | -34.1102 | -12.1882
7.0016 | -34.1106 | -12.1981
7.0077 | -34.1109 | -12.208
7.0138 | -34.1112 | -12.2179
7.0199 | -34.1115 | -12.2278
7.026 | -34.1118 | -12.2377
7.0321 | -34.1122 | -12.2474
7.0382 | -34.1125 | -12.2571
7.0443 | -34.1128 | -12.2672
7.0504 | -34.1131 | -12.2771
7.0565 | -34.1135 | -12.2871
7.0626 | -34.1138 | -12.2968
7.0687 | -34.1141 | -12.3063
7.0748 | -34.1145 | -12.316
7.0809 | -34.1148 | -12.3259
7.087 | -34.1151 | -12.3359
7.0932 | -34.1154 | -12.3459
7.0993 | -34.1157 | -12.3557
7.1054 | -34.1161 | -12.3655
7.1115 | -34.1164 | -12.3752
7.1176 | -34.1167 | -12.385
7.1237 | -34.117 | -12.3949
7.1298 | -34.1174 | -12.4047
7.1359 | -34.1177 | -12.4146
7.142 | -34.118 | -12.4242
7.1481 | -34.1183 | -12.434
7.1542 | -34.1186 | -12.4439
7.1603 | -34.119 | -12.4538
7.1664 | -34.1193 | -12.4636
7.1725 | -34.1197 | -12.4734
7.1786 | -34.12 | -12.4833
7.1847 | -34.1203 | -12.4932
7.1908 | -34.1206 | -12.5032
7.1969 | -34.1209 | -12.5132
7.203 | -34.1213 | -12.5229
7.2091 | -34.1216 | -12.5326
7.2152 | -34.1219 | -12.5425
7.2213 | -34.1223 | -12.5521
7.2274 | -34.1226 | -12.562
7.2335 | -34.1229 | -12.5717
7.2397 | -34.1233 | -12.5815
7.2458 | -34.1236 | -12.5911
7.2519 | -34.124 | -12.6009
7.258 | -34.1243 | -12.6109
7.2641 | -34.1246 | -12.6208
7.2702 | -34.1249 | -12.6308
7.2763 | -34.1253 | -12.6406
7.2824 | -34.1256 | -12.6506
7.2885 | -34.126 | -12.6604
7.2946 | -34.1263 | -12.6702
7.3007 | -34.1266 | -12.6801
7.3068 | -34.1269 | -12.6902
7.3129 | -34.1273 | -12.7
7.319 | -34.1276 | -12.7099
7.3251 | -34.1279 | -12.7195
7.3312 | -34.1283 | -12.7293
7.3373 | -34.1286 | -12.739
7.3434 | -34.129 | -12.7487
7.3495 | -34.1293 | -12.7583
7.3556 | -34.1297 | -12.7682
7.3617 | -34.13 | -12.7782
7.3678 | -34.1303 | -12.7879
7.3739 | -34.1306 | -12.7979
7.3801 | -34.131 | -12.808
7.3862 | -34.1313 | -12.8177
7.3923 | -34.1317 | -12.8277
7.3984 | -34.132 | -12.8373
7.4045 | -34.1324 | -12.847
7.4106 | -34.1327 | -12.8568
7.4167 | -34.133 | -12.8667
7.4228 | -34.1334 | -12.8766
7.4289 | -34.1337 | -12.8863
7.435 | -34.1341 | -12.8959
7.4411 | -34.1344 | -12.9057
7.4472 | -34.1347 | -12.9156
7.4533 | -34.1351 | -12.9252
7.4594 | -34.1354 | -12.9352
7.4655 | -34.1358 | -12.9448
7.4716 | -34.1361 | -12.9548
7.4777 | -34.1364 | -12.9646
7.4838 | -34.1368 | -12.9742
7.4899 | -34.1371 | -12.9839
7.496 | -34.1374 | -12.9934
7.5021 | -34.1378 | -13.0032
7.5082 | -34.1381 | -13.013
7.5143 | -34.1385 | -13.0228
7.5204 | -34.1388 | -13.0327
7.5266 | -34.1391 | -13.0423
7.5327 | -34.1395 | -13.052
7.5388 | -34.1398 | -13.062
7.5449 | -34.1401 | -13.072
7.551 | -34.1405 | -13.0816
7.5571 | -34.1408 | -13.0915
7.5632 | -34.1412 | -13.1013
7.5693 | -34.1416 | -13.111
7.5754 | -34.1419 | -13.1206
7.5815 | -34.1422 | -13.1304
7.5876 | -34.1426 | -13.1404
7.5937 | -34.1429 | -13.1502
7.5998 | -34.1433 | -13.16
7.6059 | -34.1437 | -13.1699
7.612 | -34.144 | -13.1797
7.6181 | -34.1443 | -13.1894
7.6242 | -34.1447 | -13.1994
7.6303 | -34.145 | -13.2094
7.6364 | -34.1453 | -13.2192
7.6425 | -34.1457 | -13.2291
7.6486 | -34.146 | -13.2389
7.6547 | -34.1464 | -13.2485
7.6608 | -34.1467 | -13.2583
7.667 | -34.147 | -13.2681
7.6731 | -34.1474 | -13.278
7.6792 | -34.1477 | -13.288
7.6853 | -34.1481 | -13.2979
7.6914 | -34.1484 | -13.3078
7.6975 | -34.1488 | -13.3176
7.7036 | -34.1491 | -13.3275
7.7097 | -34.1495 | -13.3374
7.7158 | -34.1498 | -13.3472
7.7219 | -34.1502 | -13.3569
7.728 | -34.1505 | -13.3667
7.7341 | -34.1509 | -13.3766
7.7402 | -34.1512 | -13.3864
7.7463 | -34.1516 | -13.3963
7.7524 | -34.1519 | -13.4063
7.7585 | -34.1523 | -13.4161
7.7646 | -34.1526 | -13.4259
7.7707 | -34.1529 | -13.4358
7.7768 | -34.1532 | -13.4458
7.7829 | -34.1536 | -13.4559
7.789 | -34.1539 | -13.4659
7.7951 | -34.1543 | -13.4757
7.8012 | -34.1546 | -13.4855
7.8073 | -34.155 | -13.4955
7.8135 | -34.1554 | -13.5053
7.8196 | -34.1557 | -13.5153
7.8257 | -34.1561 | -13.5253
7.8318 | -34.1564 | -13.5352
7.8379 | -34.1568 | -13.545
7.844 | -34.1571 | -13.5547
7.8501 | -34.1574 | -13.5648
7.8562 | -34.1578 | -13.5749
7.8623 | -34.1582 | -13.5846
7.8684 | -34.1585 | -13.5945
7.8745 | -34.1588 | -13.6044
7.8806 | -34.1591 | -13.6144
7.8867 | -34.1595 | -13.624
7.8928 | -34.1599 | -13.6337
7.8989 | -34.1602 | -13.6435
7.905 | -34.1606 | -13.6535
7.9111 | -34.1609 | -13.6634
7.9172 | -34.1612 | -13.6732
7.9233 | -34.1616 | -13.6832
7.9294 | -34.1619 | -13.6931
7.9355 | -34.1623 | -13.7028
7.9416 | -34.1626 | -13.7127
7.9477 | -34.1629 | -13.7224
7.9539 | -34.1633 | -13.7321
7.96 | -34.1636 | -13.7421
7.9661 | -34.164 | -13.7519
7.9722 | -34.1643 | -13.7616
7.9783 | -34.1647 | -13.7717
7.9844 | -34.165 | -13.7814
7.9905 | -34.1654 | -13.7915
7.9966 | -34.1657 | -13.8015
8.0027 | -34.166 | -13.8113
8.0088 | -34.1664 | -13.8211
8.0149 | -34.1667 | -13.831
8.021 | -34.1671 | -13.8409
8.0271 | -34.1674 | -13.851
8.0332 | -34.1678 | -13.8607
8.0393 | -34.1681 | -13.8707
8.0454 | -34.1685 | -13.8808
8.0515 | -34.1689 | -13.8906
8.0576 | -34.1692 | -13.9007
8.0637 | -34.1695 | -13.9105
8.0698 | -34.1699 | -13.9202
8.0759 | -34.1702 | -13.93
8.082 | -34.1705 | -13.9398
8.0881 | -34.1709 | -13.9497
8.0942 | -34.1713 | -13.9596
8.1004 | -34.1716 | -13.9695
8.1065 | -34.172 | -13.9793
8.1126 | -34.1723 | -13.9892
8.1187 | -34.1726 | -13.9992
8.1248 | -34.173 | -14.0088
8.1309 | -34.1733 | -14.0186
8.137 | -34.1737 | -14.0282
8.1431 | -34.174 | -14.0383
8.1492 | -34.1744 | -14.0481
8.1553 | -34.1747 | -14.058
8.1614 | -34.1751 | -14.0681
8.1675 | -34.1754 | -14.0779
8.1736 | -34.1757 | -14.0879
8.1797 | -34.1761 | -14.0979
8.1858 | -34.1764 | -14.1078
8.1919 | -34.1768 | -14.1176
8.198 | -34.1771 | -14.1273
8.2041 | -34.1775 | -14.1373
8.2102 | -34.1778 | -14.1471
8.2163 | -34.1782 | -14.1568
8.2224 | -34.1785 | -14.1667
8.2285 | -34.1788 | -14.1763
8.2346 | -34.1792 | -14.1861
8.2408 | -34.1795 | -14.1958
8.2469 | -34.1799 | -14.2057
8.253 | -34.1803 | -14.2155
8.2591 | -34.1806 | -14.2254
8.2652 | -34.181 | -14.2353
8.2713 | -34.1813 | -14.2452
8.2774 | -34.1817 | -14.2551
8.2835 | -34.182 | -14.2652
8.2896 | -34.1824 | -14.2749
8.2957 | -34.1827 | -14.2848
8.3018 | -34.1831 | -14.2947
8.3079 | -34.1834 | -14.3044
8.314 | -34.1838 | -14.3143
8.3201 | -34.1842 | -14.3241
8.3262 | -34.1845 | -14.334
8.3323 | -34.1849 | -14.3437
8.3384 | -34.1852 | -14.3537
8.3445 | -34.1855 | -14.3634
8.3506 | -34.1859 | -14.3733
8.3567 | -34.1863 | -14.3832
8.3628 | -34.1866 | -14.3931
8.3689 | -34.187 | -14.4028
8.375 | -34.1873 | -14.4127
8.3812 | -34.1876 | -14.4224
8.3873 | -34.188 | -14.4322
8.3934 | -34.1884 | -14.4421
8.3995 | -34.1887 | -14.4518
8.4056 | -34.189 | -14.4616
8.4117 | -34.1894 | -14.4715
8.4178 | -34.1897 | -14.4814
8.4239 | -34.1901 | -14.4911
8.43 | -34.1904 | -14.5009
8.4361 | -34.1907 | -14.5109
8.4422 | -34.1911 | -14.5206
8.4483 | -34.1914 | -14.5303
8.4544 | -34.1918 | -14.5402
8.4605 | -34.1922 | -14.5499
8.4666 | -34.1925 | -14.5598
8.4727 | -34.1928 | -14.5697
8.4788 | -34.1932 | -14.5795
8.4849 | -34.1935 | -14.5893
8.491 | -34.1939 | -14.5991
8.4971 | -34.1942 | -14.6088
8.5032 | -34.1946 | -14.6185
8.5093 | -34.1949 | -14.6281
8.5154 | -34.1953 | -14.638
8.5215 | -34.1957 | -14.6477
8.5277 | -34.196 | -14.6575
8.5338 | -34.1964 | -14.6673
8.5399 | -34.1967 | -14.677
8.546 | -34.1971 | -14.6867
8.5521 | -34.1974 | -14.6965
8.5582 | -34.1978 | -14.7063
8.5643 | -34.1981 | -14.7158
8.5704 | -34.1985 | -14.7254
8.5765 | -34.1988 | -14.7353
8.5826 | -34.1992 | -14.7447
8.5887 | -34.1996 | -14.7545
8.5948 | -34.1999 | -14.7643
8.6009 | -34.2003 | -14.7738
8.607 | -34.2006 | -14.7837
8.6131 | -34.201 | -14.7935
8.6192 | -34.2014 | -14.8031
8.6253 | -34.2017 | -14.8128
8.6314 | -34.2021 | -14.8225
8.6375 | -34.2024 | -14.8323
8.6436 | -34.2028 | -14.8421
8.6497 | -34.2031 | -14.8517
8.6558 | -34.2035 | -14.8615
8.6619 | -34.2038 | -14.8712
8.6681 | -34.2042 | -14.8812
8.6742 | -34.2046 | -14.8908
8.6803 | -34.2049 | -14.9005
8.6864 | -34.2053 | -14.91
8.6925 | -34.2056 | -14.9195
8.6986 | -34.206 | -14.9293
8.7047 | -34.2063 | -14.9392
8.7108 | -34.2067 | -14.9489
8.7169 | -34.2071 | -14.9587
8.723 | -34.2074 | -14.9683
8.7291 | -34.2078 | -14.9778
8.7352 | -34.2081 | -14.9877
8.7413 | -34.2085 | -14.9975
8.7474 | -34.2088 | -15.0071
8.7535 | -34.2092 | -15.0168
8.7596 | -34.2096 | -15.0266
8.7657 | -34.2099 | -15.0364
8.7718 | -34.2102 | -15.0462
8.7779 | -34.2106 | -15.0559
8.784 | -34.211 | -15.0657
8.7901 | -34.2114 | -15.0754
8.7962 | -34.2117 | -15.0849
8.8023 | -34.2121 | -15.0946
8.8084 | -34.2124 | -15.1042
8.8146 | -34.2128 | -15.1139
8.8207 | -34.2131 | -15.1235
8.8268 | -34.2135 | -15.133
8.8329 | -34.2139 | -15.1427
8.839 | -34.2143 | -15.1524
8.8451 | -34.2146 | -15.162
8.8512 | -34.215 | -15.1715
8.8573 | -34.2153 | -15.1813
8.8634 | -34.2157 | -15.1908
8.8695 | -34.2161 | -15.2006
8.8756 | -34.2164 | -15.2104
8.8817 | -34.2168 | -15.2199
8.8878 | -34.2171 | -15.2298
8.8939 | -34.2175 | -15.2395
8.9 | -34.2179 | -15.2491
8.9061 | -34.2182 | -15.2587
8.9122 | -34.2186 | -15.2684
8.9183 | -34.219 | -15.2781
8.9244 | -34.2194 | -15.2879
8.9305 | -34.2198 | -15.2973
8.9366 | -34.2201 | -15.307
8.9427 | -34.2205 | -15.3167
8.9488 | -34.2209 | -15.3263
8.955 | -34.2212 | -15.3359
8.9611 | -34.2216 | -15.3458
8.9672 | -34.222 | -15.3554
8.9733 | -34.2224 | -15.3649
8.9794 | -34.2228 | -15.3744
8.9855 | -34.2232 | -15.3842
8.9916 | -34.2235 | -15.394
8.9977 | -34.2239 | -15.4036
9.0038 | -34.2243 | -15.4131
9.0099 | -34.2247 | -15.4227
9.016 | -34.225 | -15.4324
9.0221 | -34.2254 | -15.4419
9.0282 | -34.2258 | -15.4513
9.0343 | -34.2262 | -15.461
9.0404 | -34.2266 | -15.4706
9.0465 | -34.2269 | -15.4804
9.0526 | -34.2273 | -15.4899
9.0587 | -34.2277 | -15.4995
9.0648 | -34.2281 | -15.509
9.0709 | -34.2285 | -15.5186
9.077 | -34.2289 | -15.5283
9.0831 | -34.2292 | -15.5381
9.0892 | -34.2296 | -15.5477
9.0953 | -34.23 | -15.5575
9.1015 | -34.2304 | -15.5671
9.1076 | -34.2308 | -15.5766
9.1137 | -34.2312 | -15.5861
9.1198 | -34.2316 | -15.5957
9.1259 | -34.232 | -15.6053
9.132 | -34.2324 | -15.6151
9.1381 | -34.2328 | -15.6246
9.1442 | -34.2332 | -15.6339
9.1503 | -34.2335 | -15.6436
9.1564 | -34.2339 | -15.6532
9.1625 | -34.2343 | -15.6629
9.1686 | -34.2347 | -15.6724
9.1747 | -34.2351 | -15.682
9.1808 | -34.2355 | -15.6916
9.1869 | -34.2359 | -15.7012
9.193 | -34.2363 | -15.7108
9.1991 | -34.2366 | -15.7205
9.2052 | -34.237 | -15.73
9.2113 | -34.2374 | -15.7397
9.2174 | -34.2379 | -15.7493
9.2235 | -34.2382 | -15.7589
9.2296 | -34.2387 | -15.7684
9.2357 | -34.239 | -15.7782
9.2419 | -34.2394 | -15.7877
9.248 | -34.2399 | -15.7973
9.2541 | -34.2403 | -15.8069
9.2602 | -34.2407 | -15.8167
9.2663 | -34.2411 | -15.826
9.2724 | -34.2415 | -15.8356
9.2785 | -34.2418 | -15.8451
9.2846 | -34.2423 | -15.8546
9.2907 | -34.2427 | -15.8642
9.2968 | -34.243 | -15.8739
9.3029 | -34.2435 | -15.8836
9.309 | -34.2438 | -15.8932
9.3151 | -34.2442 | -15.9026
9.3212 | -34.2447 | -15.912
9.3273 | -34.2451 | -15.9216
9.3334 | -34.2455 | -15.931
9.3395 | -34.2459 | -15.9405
9.3456 | -34.2462 | -15.9499
9.3517 | -34.2466 | -15.9596
9.3578 | -34.247 | -15.9691
9.3639 | -34.2474 | -15.9785
9.37 | -34.2478 | -15.9881
9.3761 | -34.2483 | -15.9978
9.3822 | -34.2487 | -16.0074
9.3884 | -34.2491 | -16.0168
9.3945 | -34.2495 | -16.0264
9.4006 | -34.2499 | -16.0359
9.4067 | -34.2503 | -16.0455
9.4128 | -34.2507 | -16.0551
9.4189 | -34.2511 | -16.0649
9.425 | -34.2515 | -16.0743
9.4311 | -34.252 | -16.0838
9.4372 | -34.2524 | -16.0933
9.4433 | -34.2528 | -16.1029
9.4494 | -34.2532 | -16.1123
9.4555 | -34.2536 | -16.1219
9.4616 | -34.254 | -16.1314
9.4677 | -34.2544 | -16.141
9.4738 | -34.2548 | -16.1505
9.4799 | -34.2552 | -16.1601
9.486 | -34.2556 | -16.1696
9.4921 | -34.256 | -16.1791
9.4982 | -34.2565 | -16.1885
9.5043 | -34.2569 | -16.198
9.5104 | -34.2573 | -16.2074
9.5165 | -34.2577 | -16.2169
9.5226 | -34.2581 | -16.2266
9.5288 | -34.2585 | -16.236
9.5349 | -34.2589 | -16.2456
9.541 | -34.2593 | -16.2549
9.5471 | -34.2597 | -16.2644
9.5532 | -34.2601 | -16.274
9.5593 | -34.2606 | -16.2837
9.5654 | -34.261 | -16.2931
9.5715 | -34.2614 | -16.3026
9.5776 | -34.2618 | -16.3123
9.5837 | -34.2622 | -16.3217
9.5898 | -34.2626 | -16.3313
9.5959 | -34.263 | -16.3409
9.602 | -34.2634 | -16.3502
9.6081 | -34.2638 | -16.3597
9.6142 | -34.2642 | -16.3692
9.6203 | -34.2646 | -16.3787
9.6264 | -34.265 | -16.3881
9.6325 | -34.2654 | -16.3977
9.6386 | -34.2658 | -16.4071
9.6447 | -34.2662 | -16.4167
9.6508 | -34.2666 | -16.4261
9.6569 | -34.267 | -16.4356
9.663 | -34.2675 | -16.4452
9.6691 | -34.2679 | -16.4547
9.6753 | -34.2683 | -16.4641
9.6814 | -34.2687 | -16.4735
9.6875 | -34.2691 | -16.4828
9.6936 | -34.2695 | -16.4921
9.6997 | -34.27 | -16.5017
9.7058 | -34.2704 | -16.5112
9.7119 | -34.2708 | -16.5208
9.718 | -34.2712 | -16.5304
9.7241 | -34.2717 | -16.5398
9.7302 | -34.2721 | -16.5493
9.7363 | -34.2725 | -16.5587
9.7424 | -34.2729 | -16.568
9.7485 | -34.2733 | -16.5775
9.7546 | -34.2737 | -16.5869
9.7607 | -34.2741 | -16.5965
9.7668 | -34.2745 | -16.6061
9.7729 | -34.275 | -16.6155
9.779 | -34.2754 | -16.6247
9.7851 | -34.2758 | -16.6344
9.7912 | -34.2763 | -16.6438
9.7973 | -34.2767 | -16.6533
9.8034 | -34.2771 | -16.6626
9.8095 | -34.2775 | -16.6722
9.8157 | -34.278 | -16.6817
9.8218 | -34.2784 | -16.6913
9.8279 | -34.2788 | -16.7006
9.834 | -34.2792 | -16.7102
9.8401 | -34.2796 | -16.7197
9.8462 | -34.28 | -16.7289
9.8523 | -34.2804 | -16.7383
9.8584 | -34.2808 | -16.7478
9.8645 | -34.2813 | -16.7574
9.8706 | -34.2816 | -16.7669
9.8767 | -34.282 | -16.7763
9.8828 | -34.2824 | -16.7857
9.8889 | -34.2829 | -16.7949
9.895 | -34.2833 | -16.8044
9.9011 | -34.2837 | -16.8139
9.9072 | -34.2841 | -16.8233
9.9133 | -34.2845 | -16.8327
9.9194 | -34.2849 | -16.842
9.9255 | -34.2853 | -16.8514
9.9316 | -34.2857 | -16.8608
9.9377 | -34.2862 | -16.8702
9.9438 | -34.2866 | -16.8797
9.9499 | -34.287 | -16.8893
9.956 | -34.2874 | -16.8987
9.9622 | -34.2878 | -16.9083
9.9683 | -34.2882 | -16.9177
9.9744 | -34.2886 | -16.9273
9.9805 | -34.2891 | -16.9367
9.9866 | -34.2895 | -16.9461
9.9927 | -34.2899 | -16.9554
9.9988 | -34.2903 | -16.965
10.0049 | -34.2907 | -16.9744
10.011 | -34.2911 | -16.9838
10.0171 | -34.2915 | -16.9932
10.0232 | -34.292 | -17.0025
10.0293 | -34.2924 | -17.0119
10.0354 | -34.2928 | -17.0214
10.0415 | -34.2933 | -17.0308
10.0476 | -34.2937 | -17.0404
10.0537 | -34.2941 | -17.0499
10.0598 | -34.2945 | -17.0593
10.0659 | -34.2949 | -17.0688
10.072 | -34.2953 | -17.078
10.0781 | -34.2957 | -17.0874
10.0842 | -34.2961 | -17.0968
10.0903 | -34.2966 | -17.1062
10.0964 | -34.297 | -17.1153
10.1026 | -34.2974 | -17.1248
10.1087 | -34.2978 | -17.1342
10.1148 | -34.2982 | -17.1436
10.1209 | -34.2986 | -17.1529
10.127 | -34.299 | -17.1622
10.1331 | -34.2994 | -17.1715
10.1392 | -34.2998 | -17.1807
10.1453 | -34.3002 | -17.1901
10.1514 | -34.3007 | -17.1995
10.1575 | -34.301 | -17.209
10.1636 | -34.3015 | -17.2181
10.1697 | -34.3019 | -17.2276
10.1758 | -34.3023 | -17.2368
10.1819 | -34.3027 | -17.2461
10.188 | -34.3031 | -17.2556
10.1941 | -34.3035 | -17.265
10.2002 | -34.3039 | -17.2743
10.2063 | -34.3043 | -17.2837
10.2124 | -34.3047 | -17.2931
10.2185 | -34.3052 | -17.3025
10.2246 | -34.3056 | -17.3118
10.2307 | -34.306 | -17.3212
10.2368 | -34.3064 | -17.3303
10.2429 | -34.3068 | -17.3397
10.2491 | -34.3072 | -17.3491
10.2552 | -34.3076 | -17.3585
10.2613 | -34.308 | -17.3679
10.2674 | -34.3084 | -17.3773
10.2735 | -34.3089 | -17.3868
10.2796 | -34.3093 | -17.3962
10.2857 | -34.3097 | -17.4057
10.2918 | -34.3101 | -17.4152
10.2979 | -34.3105 | -17.4245
10.304 | -34.311 | -17.4337
10.3101 | -34.3113 | -17.4431
10.3162 | -34.3118 | -17.4526
10.3223 | -34.3122 | -17.4618
10.3284 | -34.3126 | -17.4711
10.3345 | -34.313 | -17.4804
10.3406 | -34.3134 | -17.4897
10.3467 | -34.3138 | -17.4989
10.3528 | -34.3142 | -17.5082
10.3589 | -34.3146 | -17.5176
10.365 | -34.315 | -17.527
10.3711 | -34.3154 | -17.5365
10.3772 | -34.3159 | -17.5458
10.3833 | -34.3163 | -17.5551
10.3895 | -34.3167 | -17.5645
10.3956 | -34.3171 | -17.5738
10.4017 | -34.3175 | -17.583
10.4078 | -34.3179 | -17.5923
10.4139 | -34.3183 | -17.6017
10.42 | -34.3188 | -17.6111
10.4261 | -34.3192 | -17.6204
10.4322 | -34.3196 | -17.6297
10.4383 | -34.32 | -17.6391
10.4444 | -34.3204 | -17.6482
10.4505 | -34.3208 | -17.6576
10.4566 | -34.3212 | -17.6668
10.4627 | -34.3216 | -17.676
10.4688 | -34.322 | -17.6852
10.4749 | -34.3224 | -17.6946
10.481 | -34.3228 | -17.704
10.4871 | -34.3232 | -17.7132
10.4932 | -34.3237 | -17.7224
10.4993 | -34.3241 | -17.7317
10.5054 | -34.3245 | -17.7411
10.5115 | -34.3249 | -17.7504
10.5176 | -34.3253 | -17.7595
10.5237 | -34.3258 | -17.769
10.5298 | -34.3262 | -17.7782
10.536 | -34.3266 | -17.7875
10.5421 | -34.327 | -17.7967
10.5482 | -34.3274 | -17.8061
10.5543 | -34.3278 | -17.8152
10.5604 | -34.3282 | -17.8244
10.5665 | -34.3286 | -17.8335
10.5726 | -34.329 | -17.8428
10.5787 | -34.3294 | -17.852
10.5848 | -34.3299 | -17.861
10.5909 | -34.3303 | -17.8704
10.597 | -34.3307 | -17.8797
10.6031 | -34.3311 | -17.8892
10.6092 | -34.3315 | -17.8984
10.6153 | -34.3319 | -17.9078
10.6214 | -34.3323 | -17.9169
10.6275 | -34.3327 | -17.9261
10.6336 | -34.3332 | -17.9354
10.6397 | -34.3335 | -17.9447
10.6458 | -34.334 | -17.954
10.6519 | -34.3344 | -17.9633
10.658 | -34.3348 | -17.9725
10.6641 | -34.3352 | -17.9819
10.6702 | -34.3356 | -17.9912
10.6764 | -34.336 | -18.0005
10.6825 | -34.3364 | -18.0097
10.6886 | -34.3368 | -18.0188
10.6947 | -34.3373 | -18.028
10.7008 | -34.3377 | -18.0373
10.7069 | -34.3381 | -18.0466
10.713 | -34.3385 | -18.0559
10.7191 | -34.3389 | -18.0651
10.7252 | -34.3393 | -18.0741
10.7313 | -34.3398 | -18.0833
10.7374 | -34.3401 | -18.0925
10.7435 | -34.3406 | -18.1018
10.7496 | -34.341 | -18.1109
10.7557 | -34.3414 | -18.1202
10.7618 | -34.3418 | -18.1292
10.7679 | -34.3422 | -18.1384
10.774 | -34.3426 | -18.1478
10.7801 | -34.343 | -18.157
10.7862 | -34.3434 | -18.166
10.7923 | -34.3438 | -18.1754
10.7984 | -34.3443 | -18.1844
10.8045 | -34.3447 | -18.1935
10.8106 | -34.3451 | -18.2028
10.8168 | -34.3455 | -18.2119
10.8229 | -34.3459 | -18.2211
10.829 | -34.3463 | -18.2302
10.8351 | -34.3467 | -18.2392
10.8412 | -34.3472 | -18.2483
10.8473 | -34.3476 | -18.2573
10.8534 | -34.348 | -18.2665
10.8595 | -34.3484 | -18.2757
10.8656 | -34.3488 | -18.2847
10.8717 | -34.3492 | -18.2938
10.8778 | -34.3497 | -18.3031
10.8839 | -34.3501 | -18.3122
10.89 | -34.3505 | -18.3214
10.8961 | -34.3509 | -18.3306
10.9022 | -34.3513 | -18.3397
10.9083 | -34.3517 | -18.3488
10.9144 | -34.3521 | -18.3579
10.9205 | -34.3526 | -18.367
10.9266 | -34.353 | -18.3762
10.9327 | -34.3534 | -18.3853
10.9388 | -34.3538 | -18.3943
10.9449 | -34.3542 | -18.4034
10.951 | -34.3547 | -18.4126
10.9571 | -34.3551 | -18.4217
10.9633 | -34.3555 | -18.431
10.9694 | -34.3559 | -18.4403
10.9755 | -34.3563 | -18.4494
10.9816 | -34.3567 | -18.4586
10.9877 | -34.3572 | -18.4677
10.9938 | -34.3576 | -18.477
10.9999 | -34.358 | -18.486
11.006 | -34.3584 | -18.495
11.0121 | -34.3588 | -18.5041
11.0182 | -34.3593 | -18.5132
11.0243 | -34.3597 | -18.5224
11.0304 | -34.3601 | -18.5315
11.0365 | -34.3605 | -18.5405
11.0426 | -34.3609 | -18.5496
11.0487 | -34.3614 | -18.5588
11.0548 | -34.3618 | -18.5678
11.0609 | -34.3622 | -18.577
11.067 | -34.3626 | -18.5861
11.0731 | -34.3631 | -18.5951
11.0792 | -34.3635 | -18.6041
11.0853 | -34.3639 | -18.6133
11.0914 | -34.3643 | -18.6225
11.0975 | -34.3648 | -18.6315
11.1037 | -34.3652 | -18.6406
11.1098 | -34.3656 | -18.6499
11.1159 | -34.366 | -18.6587
11.122 | -34.3664 | -18.6677
11.1281 | -34.3668 | -18.6769
11.1342 | -34.3673 | -18.686
11.1403 | -34.3677 | -18.6952
11.1464 | -34.3681 | -18.7044
11.1525 | -34.3685 | -18.7133
11.1586 | -34.3689 | -18.7225
11.1647 | -34.3694 | -18.7315
11.1708 | -34.3698 | -18.7406
11.1769 | -34.3702 | -18.7496
11.183 | -34.3706 | -18.7586
11.1891 | -34.3711 | -18.7677
11.1952 | -34.3715 | -18.7767
11.2013 | -34.3719 | -18.7859
11.2074 | -34.3723 | -18.7949
11.2135 | -34.3728 | -18.8042
11.2196 | -34.3732 | -18.8131
11.2257 | -34.3736 | -18.8222
11.2318 | -34.3741 | -18.8312
11.2379 | -34.3745 | -18.8403
11.244 | -34.3749 | -18.8495
11.2502 | -34.3754 | -18.8585
11.2563 | -34.3758 | -18.8675
11.2624 | -34.3762 | -18.8765
11.2685 | -34.3767 | -18.8855
11.2746 | -34.3771 | -18.8944
11.2807 | -34.3775 | -18.9037
11.2868 | -34.3779 | -18.9128
11.2929 | -34.3784 | -18.9216
11.299 | -34.3788 | -18.9307
11.3051 | -34.3792 | -18.9399
11.3112 | -34.3797 | -18.9491
11.3173 | -34.3801 | -18.958
11.3234 | -34.3806 | -18.9668
11.3295 | -34.381 | -18.976
11.3356 | -34.3814 | -18.9852
11.3417 | -34.3818 | -18.9941
11.3478 | -34.3822 | -19.0033
11.3539 | -34.3826 | -19.0124
11.36 | -34.3831 | -19.0213
11.3661 | -34.3835 | -19.0303
11.3722 | -34.3839 | -19.0392
11.3783 | -34.3844 | -19.0483
11.3844 | -34.3848 | -19.0573
11.3906 | -34.3852 | -19.0663
11.3967 | -34.3857 | -19.0752
11.4028 | -34.3861 | -19.0843
11.4089 | -34.3865 | -19.0932
11.415 | -34.3869 | -19.1023
11.4211 | -34.3873 | -19.1114
11.4272 | -34.3878 | -19.1205
11.4333 | -34.3882 | -19.1293
11.4394 | -34.3886 | -19.1383
11.4455 | -34.389 | -19.1472
11.4516 | -34.3894 | -19.1563
11.4577 | -34.3899 | -19.1653
11.4638 | -34.3903 | -19.1741
11.4699 | -34.3907 | -19.1832
11.476 | -34.3912 | -19.1923
11.4821 | -34.3916 | -19.2013
11.4882 | -34.392 | -19.2104
11.4943 | -34.3924 | -19.2195
11.5004 | -34.3929 | -19.2283
11.5065 | -34.3933 | -19.2372
11.5126 | -34.3938 | -19.2461
11.5187 | -34.3942 | -19.2551
11.5248 | -34.3946 | -19.2643
11.5309 | -34.3951 | -19.2732
11.5371 | -34.3955 | -19.2821
11.5432 | -34.3959 | -19.2912
11.5493 | -34.3963 | -19.3003
11.5554 | -34.3968 | -19.3093
11.5615 | -34.3972 | -19.3184
11.5676 | -34.3976 | -19.3274
11.5737 | -34.3981 | -19.3362
11.5798 | -34.3985 | -19.345
11.5859 | -34.399 | -19.354
11.592 | -34.3994 | -19.3629
11.5981 | -34.3998 | -19.3718
11.6042 | -34.4002 | -19.3808
11.6103 | -34.4006 | -19.3896
11.6164 | -34.4011 | -19.3985
11.6225 | -34.4015 | -19.4074
11.6286 | -34.4019 | -19.4163
11.6347 | -34.4024 | -19.4254
11.6408 | -34.4028 | -19.4342
11.6469 | -34.4032 | -19.4431
11.653 | -34.4037 | -19.452
11.6591 | -34.4041 | -19.4609
11.6652 | -34.4045 | -19.4698
11.6713 | -34.405 | -19.4787
11.6775 | -34.4054 | -19.4875
11.6836 | -34.4058 | -19.4964
11.6897 | -34.4063 | -19.5053
11.6958 | -34.4067 | -19.5143
11.7019 | -34.4071 | -19.5234
11.708 | -34.4076 | -19.5324
11.7141 | -34.408 | -19.5411
11.7202 | -34.4085 | -19.5501
11.7263 | -34.4089 | -19.559
11.7324 | -34.4093 | -19.5681
11.7385 | -34.4097 | -19.5771
11.7446 | -34.4101 | -19.5859
11.7507 | -34.4106 | -19.5947
11.7568 | -34.411 | -19.6036
11.7629 | -34.4114 | -19.6126
11.769 | -34.4119 | -19.6214
11.7751 | -34.4123 | -19.6303
11.7812 | -34.4128 | -19.639
11.7873 | -34.4132 | -19.648
11.7934 | -34.4136 | -19.6569
11.7995 | -34.4141 | -19.6658
11.8056 | -34.4145 | -19.6746
11.8117 | -34.4149 | -19.6835
11.8178 | -34.4154 | -19.6923
11.824 | -34.4158 | -19.7011
11.8301 | -34.4162 | -19.7099
11.8362 | -34.4167 | -19.7187
11.8423 | -34.4171 | -19.7276
11.8484 | -34.4175 | -19.7365
11.8545 | -34.418 | -19.7452
11.8606 | -34.4184 | -19.754
11.8667 | -34.4188 | -19.763
11.8728 | -34.4193 | -19.7717
11.8789 | -34.4197 | -19.7806
11.885 | -34.4201 | -19.7894
11.8911 | -34.4206 | -19.7983
11.8972 | -34.421 | -19.807
11.9033 | -34.4215 | -19.8159
11.9094 | -34.4219 | -19.8247
11.9155 | -34.4223 | -19.8335
11.9216 | -34.4228 | -19.8422
11.9277 | -34.4232 | -19.851
11.9338 | -34.4236 | -19.8599
11.9399 | -34.4241 | -19.8689
11.946 | -34.4245 | -19.8777
11.9521 | -34.4249 | -19.8865
11.9582 | -34.4254 | -19.8952
11.9644 | -34.4258 | -19.904
11.9705 | -34.4263 | -19.9128
11.9766 | -34.4267 | -19.9216
11.9827 | -34.4272 | -19.9304
11.9888 | -34.4276 | -19.9392
11.9949 | -34.4281 | -19.9477
12.001 | -34.4285 | -19.9566
12.0071 | -34.4289 | -19.9653
12.0132 | -34.4294 | -19.9741
12.0193 | -34.4298 | -19.983
12.0254 | -34.4302 | -19.9918
12.0315 | -34.4307 | -20.0004
12.0376 | -34.4311 | -20.0091
12.0437 | -34.4315 | -20.0179
12.0498 | -34.432 | -20.0268
12.0559 | -34.4324 | -20.0355
12.062 | -34.4329 | -20.0441
12.0681 | -34.4333 | -20.0529
12.0742 | -34.4338 | -20.0616
12.0803 | -34.4342 | -20.0705
12.0864 | -34.4346 | -20.0794
12.0925 | -34.4351 | -20.0882
12.0986 | -34.4355 | -20.0969
12.1047 | -34.436 | -20.1057
12.1109 | -34.4364 | -20.1143
12.117 | -34.4369 | -20.123
12.1231 | -34.4373 | -20.1317
12.1292 | -34.4377 | -20.1405
12.1353 | -34.4382 | -20.1491
12.1414 | -34.4386 | -20.158
12.1475 | -34.4391 | -20.1668
12.1536 | -34.4395 | -20.1756
12.1597 | -34.4399 | -20.1844
12.1658 | -34.4404 | -20.1931
12.1719 | -34.4408 | -20.2016
12.178 | -34.4413 | -20.2104
12.1841 | -34.4417 | -20.2192
12.1902 | -34.4422 | -20.228
12.1963 | -34.4426 | -20.2368
12.2024 | -34.4431 | -20.2453
12.2085 | -34.4435 | -20.2541
12.2146 | -34.444 | -20.2629
12.2207 | -34.4444 | -20.2716
12.2268 | -34.4449 | -20.2803
12.2329 | -34.4453 | -20.289
12.239 | -34.4458 | -20.2975
12.2451 | -34.4462 | -20.3062
12.2513 | -34.4466 | -20.3151
12.2574 | -34.4471 | -20.3238
12.2635 | -34.4475 | -20.3325
12.2696 | -34.448 | -20.3411
12.2757 | -34.4484 | -20.3496
12.2818 | -34.4489 | -20.3583
12.2879 | -34.4493 | -20.3671
12.294 | -34.4498 | -20.3756
12.3001 | -34.4502 | -20.3843
12.3062 | -34.4507 | -20.393
12.3123 | -34.4511 | -20.4014
12.3184 | -34.4516 | -20.41
12.3245 | -34.452 | -20.4186
12.3306 | -34.4524 | -20.4273
12.3367 | -34.4529 | -20.436
12.3428 | -34.4533 | -20.4447
12.3489 | -34.4538 | -20.453
12.355 | -34.4543 | -20.4618
12.3611 | -34.4547 | -20.4706
12.3672 | -34.4551 | -20.4792
12.3733 | -34.4556 | -20.4879
12.3794 | -34.456 | -20.4963
12.3855 | -34.4565 | -20.5049
12.3916 | -34.4569 | -20.5136
12.3978 | -34.4574 | -20.5223
12.4039 | -34.4578 | -20.531
12.41 | -34.4583 | -20.5397
12.4161 | -34.4588 | -20.5481
12.4222 | -34.4592 | -20.5567
12.4283 | -34.4596 | -20.5653
12.4344 | -34.4601 | -20.5741
12.4405 | -34.4605 | -20.5828
12.4466 | -34.461 | -20.5913
12.4527 | -34.4614 | -20.5999
12.4588 | -34.4619 | -20.6085
12.4649 | -34.4623 | -20.6172
12.471 | -34.4628 | -20.6259
12.4771 | -34.4632 | -20.6345
12.4832 | -34.4637 | -20.6431
12.4893 | -34.4641 | -20.6515
12.4954 | -34.4646 | -20.66
12.5015 | -34.465 | -20.6686
12.5076 | -34.4655 | -20.6773
12.5137 | -34.4659 | -20.6859
12.5198 | -34.4664 | -20.6944
12.5259 | -34.4668 | -20.7031
12.532 | -34.4672 | -20.7117
12.5382 | -34.4677 | -20.7203
12.5443 | -34.4681 | -20.7289
12.5504 | -34.4686 | -20.7376
12.5565 | -34.469 | -20.7461
12.5626 | -34.4695 | -20.7548
12.5687 | -34.4699 | -20.7635
12.5748 | -34.4704 | -20.7721
12.5809 | -34.4708 | -20.7808
12.587 | -34.4713 | -20.7894
12.5931 | -34.4718 | -20.7978
12.5992 | -34.4722 | -20.8065
12.6053 | -34.4727 | -20.815
12.6114 | -34.4731 | -20.8235
12.6175 | -34.4736 | -20.8322
12.6236 | -34.474 | -20.8408
12.6297 | -34.4744 | -20.8493
12.6358 | -34.4749 | -20.8578
12.6419 | -34.4753 | -20.8664
12.648 | -34.4758 | -20.875
12.6541 | -34.4762 | -20.8836
12.6602 | -34.4767 | -20.8919
12.6663 | -34.4772 | -20.9005
12.6724 | -34.4776 | -20.909
12.6785 | -34.4781 | -20.9176
12.6847 | -34.4785 | -20.926
12.6908 | -34.479 | -20.9346
12.6969 | -34.4794 | -20.9431
12.703 | -34.4799 | -20.9517
12.7091 | -34.4803 | -20.9603
12.7152 | -34.4808 | -20.9687
12.7213 | -34.4812 | -20.9773
12.7274 | -34.4817 | -20.9859
12.7335 | -34.4821 | -20.9942
12.7396 | -34.4826 | -21.0026
12.7457 | -34.483 | -21.0113
12.7518 | -34.4834 | -21.0198
12.7579 | -34.4839 | -21.0285
12.764 | -34.4843 | -21.0371
12.7701 | -34.4848 | -21.0456
12.7762 | -34.4853 | -21.054
12.7823 | -34.4857 | -21.0626
12.7884 | -34.4862 | -21.0712
12.7945 | -34.4866 | -21.0799
12.8006 | -34.4871 | -21.0884
12.8067 | -34.4875 | -21.0969
12.8128 | -34.488 | -21.1053
12.8189 | -34.4884 | -21.1138
12.8251 | -34.4889 | -21.1223
12.8312 | -34.4893 | -21.131
12.8373 | -34.4898 | -21.1395
12.8434 | -34.4902 | -21.148
12.8495 | -34.4907 | -21.1565
12.8556 | -34.4911 | -21.1649
12.8617 | -34.4916 | -21.1736
12.8678 | -34.492 | -21.1822
12.8739 | -34.4925 | -21.1906
12.88 | -34.4929 | -21.1991
12.8861 | -34.4933 | -21.2076
12.8922 | -34.4938 | -21.2162
12.8983 | -34.4943 | -21.2246
12.9044 | -34.4947 | -21.233
12.9105 | -34.4952 | -21.2415
12.9166 | -34.4956 | -21.25
12.9227 | -34.4961 | -21.2585
12.9288 | -34.4965 | -21.2671
12.9349 | -34.497 | -21.2755
12.941 | -34.4974 | -21.284
12.9471 | -34.4979 | -21.2925
12.9532 | -34.4984 | -21.301
12.9593 | -34.4988 | -21.3096
12.9654 | -34.4993 | -21.3182
12.9716 | -34.4997 | -21.3269
12.9777 | -34.5002 | -21.3353
12.9838 | -34.5006 | -21.3438
12.9899 | -34.5011 | -21.3523
12.996 | -34.5015 | -21.3608
13.0021 | -34.502 | -21.3693
13.0082 | -34.5024 | -21.3777
13.0143 | -34.5029 | -21.3861
13.0204 | -34.5034 | -21.3944
13.0265 | -34.5038 | -21.4028
13.0326 | -34.5043 | -21.4114
13.0387 | -34.5047 | -21.42
13.0448 | -34.5052 | -21.4285
13.0509 | -34.5056 | -21.4369
13.057 | -34.5061 | -21.4453
13.0631 | -34.5066 | -21.4537
13.0692 | -34.507 | -21.4621
13.0753 | -34.5075 | -21.4707
13.0814 | -34.5079 | -21.4794
13.0875 | -34.5084 | -21.4877
13.0936 | -34.5088 | -21.4963
13.0997 | -34.5093 | -21.5046
13.1058 | -34.5097 | -21.513
13.112 | -34.5102 | -21.5216
13.1181 | -34.5107 | -21.53
13.1242 | -34.5111 | -21.5384
13.1303 | -34.5116 | -21.5469
13.1364 | -34.512 | -21.5554
13.1425 | -34.5124 | -21.5637
13.1486 | -34.5129 | -21.5722
13.1547 | -34.5133 | -21.5806
13.1608 | -34.5138 | -21.5891
13.1669 | -34.5142 | -21.5974
13.173 | -34.5147 | -21.6059
13.1791 | -34.5151 | -21.6142
13.1852 | -34.5156 | -21.6227
13.1913 | -34.5161 | -21.6311
13.1974 | -34.5165 | -21.6396
13.2035 | -34.5169 | -21.6481
13.2096 | -34.5174 | -21.6565
13.2157 | -34.5179 | -21.6648
13.2218 | -34.5183 | -21.6732
13.2279 | -34.5188 | -21.6816
13.234 | -34.5192 | -21.69
13.2401 | -34.5197 | -21.6985
13.2462 | -34.5201 | -21.7069
13.2524 | -34.5206 | -21.7154
13.2585 | -34.521 | -21.7236
13.2646 | -34.5215 | -21.732
13.2707 | -34.522 | -21.7405
13.2768 | -34.5224 | -21.749
13.2829 | -34.5229 | -21.7575
13.289 | -34.5233 | -21.7659
13.2951 | -34.5238 | -21.7743
13.3012 | -34.5243 | -21.7827
13.3073 | -34.5247 | -21.7912
13.3134 | -34.5252 | -21.7997
13.3195 | -34.5256 | -21.8081
13.3256 | -34.5261 | -21.8163
13.3317 | -34.5265 | -21.8246
13.3378 | -34.527 | -21.833
13.3439 | -34.5274 | -21.8415
13.35 | -34.5279 | -21.85
13.3561 | -34.5283 | -21.8584
13.3622 | -34.5288 | -21.8667
13.3683 | -34.5292 | -21.8749
13.3744 | -34.5297 | -21.8832
13.3805 | -34.5302 | -21.8918
13.3866 | -34.5306 | -21.9001
13.3927 | -34.5311 | -21.9087
13.3989 | -34.5315 | -21.9169
13.405 | -34.532 | -21.9254
13.4111 | -34.5324 | -21.9336
13.4172 | -34.5329 | -21.942
13.4233 | -34.5333 | -21.9506
13.4294 | -34.5337 | -21.9591
13.4355 | -34.5342 | -21.9672
13.4416 | -34.5347 | -21.9755
13.4477 | -34.5352 | -21.9839
13.4538 | -34.5356 | -21.9924
13.4599 | -34.5361 | -22.0008
13.466 | -34.5365 | -22.0091
13.4721 | -34.537 | -22.0175
13.4782 | -34.5375 | -22.0258
13.4843 | -34.5379 | -22.0343
13.4904 | -34.5384 | -22.0428
13.4965 | -34.5388 | -22.0511
13.5026 | -34.5393 | -22.0596
13.5087 | -34.5397 | -22.0677
13.5148 | -34.5402 | -22.0762
13.5209 | -34.5406 | -22.0847
13.527 | -34.5411 | -22.0931
13.5331 | -34.5415 | -22.1014
13.5393 | -34.542 | -22.1097
13.5454 | -34.5424 | -22.118
13.5515 | -34.5429 | -22.1265
13.5576 | -34.5433 | -22.1348
13.5637 | -34.5438 | -22.1432
13.5698 | -34.5442 | -22.1515
13.5759 | -34.5447 | -22.1597
13.582 | -34.5451 | -22.168
13.5881 | -34.5456 | -22.1764
13.5942 | -34.546 | -22.1849
13.6003 | -34.5465 | -22.1932
13.6064 | -34.547 | -22.2016
13.6125 | -34.5474 | -22.2098
13.6186 | -34.5479 | -22.2182
13.6247 | -34.5484 | -22.2265
13.6308 | -34.5488 | -22.2349
13.6369 | -34.5492 | -22.2433
13.643 | -34.5497 | -22.2518
13.6491 | -34.5501 | -22.2601
13.6552 | -34.5506 | -22.2685
13.6613 | -34.551 | -22.2768
13.6674 | -34.5515 | -22.2852
13.6735 | -34.5519 | -22.2935
13.6796 | -34.5524 | -22.3018
13.6858 | -34.5528 | -22.3101
13.6919 | -34.5533 | -22.3184
13.698 | -34.5537 | -22.3267
13.7041 | -34.5542 | -22.3352
13.7102 | -34.5547 | -22.3435
13.7163 | -34.5551 | -22.3519
13.7224 | -34.5556 | -22.3603
13.7285 | -34.556 | -22.3687
13.7346 | -34.5565 | -22.377
13.7407 | -34.5569 | -22.3853
13.7468 | -34.5574 | -22.3937
13.7529 | -34.5578 | -22.402
13.759 | -34.5583 | -22.4106
13.7651 | -34.5587 | -22.419
13.7712 | -34.5592 | -22.4273
13.7773 | -34.5596 | -22.4357
13.7834 | -34.5601 | -22.444
13.7895 | -34.5606 | -22.4522
13.7956 | -34.561 | -22.4608
13.8017 | -34.5614 | -22.4691
13.8078 | -34.5619 | -22.4776
13.8139 | -34.5623 | -22.4859
13.82 | -34.5628 | -22.4942
13.8262 | -34.5632 | -22.5024
13.8323 | -34.5637 | -22.5109
13.8384 | -34.5642 | -22.5193
13.8445 | -34.5647 | -22.5276
13.8506 | -34.5651 | -22.5359
13.8567 | -34.5656 | -22.5443
13.8628 | -34.566 | -22.5526
13.8689 | -34.5665 | -22.561
13.875 | -34.567 | -22.5693
13.8811 | -34.5674 | -22.5777
13.8872 | -34.5678 | -22.5861
13.8933 | -34.5683 | -22.5942
13.8994 | -34.5687 | -22.6026
13.9055 | -34.5692 | -22.6111
13.9116 | -34.5696 | -22.6194
13.9177 | -34.5701 | -22.6278
13.9238 | -34.5705 | -22.6363
13.9299 | -34.571 | -22.6445
13.936 | -34.5714 | -22.653
13.9421 | -34.5719 | -22.6613
13.9482 | -34.5723 | -22.6695
13.9543 | -34.5728 | -22.6779
13.9604 | -34.5732 | -22.6861
13.9665 | -34.5737 | -22.6944
13.9727 | -34.5742 | -22.7029
13.9788 | -34.5746 | -22.7111
13.9849 | -34.5751 | -22.7195
13.991 | -34.5755 | -22.7278
13.9971 | -34.576 | -22.736
14.0032 | -34.5764 | -22.7443
14.0093 | -34.5769 | -22.7526
14.0154 | -34.5773 | -22.7608
14.0215 | -34.5778 | -22.7691
14.0276 | -34.5782 | -22.7775
14.0337 | -34.5787 | -22.7858
14.0398 | -34.5792 | -22.7941
14.0459 | -34.5796 | -22.8025
14.052 | -34.5801 | -22.8109
14.0581 | -34.5805 | -22.8192
14.0642 | -34.581 | -22.8275
14.0703 | -34.5814 | -22.8358
14.0764 | -34.5819 | -22.8439
14.0825 | -34.5824 | -22.8522
14.0886 | -34.5828 | -22.8605
14.0947 | -34.5833 | -22.8689
14.1008 | -34.5837 | -22.8772
14.1069 | -34.5842 | -22.8854
14.1131 | -34.5846 | -22.8938
14.1192 | -34.5851 | -22.9019
14.1253 | -34.5855 | -22.9101
14.1314 | -34.586 | -22.9184
14.1375 | -34.5865 | -22.9266
14.1436 | -34.5869 | -22.9348
14.1497 | -34.5874 | -22.943
14.1558 | -34.5878 | -22.9513
14.1619 | -34.5883 | -22.9596
14.168 | -34.5887 | -22.9678
14.1741 | -34.5892 | -22.976
14.1802 | -34.5897 | -22.9842
14.1863 | -34.5901 | -22.9926
14.1924 | -34.5906 | -23.0009
14.1985 | -34.5911 | -23.0091
14.2046 | -34.5915 | -23.0175
14.2107 | -34.592 | -23.0255
14.2168 | -34.5924 | -23.0338
14.2229 | -34.5929 | -23.0421
14.229 | -34.5933 | -23.0504
14.2351 | -34.5938 | -23.0587
14.2412 | -34.5942 | -23.067
14.2473 | -34.5947 | -23.0752
14.2534 | -34.5951 | -23.0834
14.2596 | -34.5956 | -23.0917
14.2657 | -34.596 | -23.0999
14.2718 | -34.5965 | -23.1083
14.2779 | -34.5969 | -23.1166
14.284 | -34.5974 | -23.1248
14.2901 | -34.5979 | -23.1331
14.2962 | -34.5984 | -23.1413
14.3023 | -34.5988 | -23.1496
14.3084 | -34.5992 | -23.1578
14.3145 | -34.5997 | -23.1661
14.3206 | -34.6002 | -23.1745
14.3267 | -34.6006 | -23.1828
14.3328 | -34.6011 | -23.191
14.3389 | -34.6015 | -23.1991
14.345 | -34.602 | -23.2074
14.3511 | -34.6024 | -23.2157
14.3572 | -34.6029 | -23.2241
14.3633 | -34.6033 | -23.2323
14.3694 | -34.6038 | -23.2405
14.3755 | -34.6043 | -23.2487
14.3816 | -34.6047 | -23.2571
14.3877 | -34.6052 | -23.2653
14.3938 | -34.6057 | -23.2737
14.4 | -34.6061 | -23.2818
14.4061 | -34.6066 | -23.29
14.4122 | -34.607 | -23.2983
14.4183 | -34.6075 | -23.3064
14.4244 | -34.608 | -23.3146
14.4305 | -34.6084 | -23.3229
14.4366 | -34.6089 | -23.3311
14.4427 | -34.6094 | -23.3395
14.4488 | -34.6098 | -23.348
14.4549 | -34.6103 | -23.3561
14.461 | -34.6108 | -23.3642
14.4671 | -34.6112 | -23.3726
14.4732 | -34.6117 | -23.3809
14.4793 | -34.6121 | -23.3892
14.4854 | -34.6126 | -23.3974
14.4915 | -34.6131 | -23.4053
14.4976 | -34.6135 | -23.4134
14.5037 | -34.614 | -23.4215
14.5098 | -34.6145 | -23.4297
14.5159 | -34.6149 | -23.438
14.522 | -34.6154 | -23.4462
14.5281 | -34.6158 | -23.4544
14.5342 | -34.6163 | -23.4628
14.5403 | -34.6168 | -23.4708
14.5465 | -34.6172 | -23.4791
14.5526 | -34.6177 | -23.4875
14.5587 | -34.6181 | -23.4957
14.5648 | -34.6186 | -23.5038
14.5709 | -34.6191 | -23.5119
14.577 | -34.6195 | -23.5201
14.5831 | -34.6199 | -23.5284
14.5892 | -34.6204 | -23.5365
14.5953 | -34.6209 | -23.5445
14.6014 | -34.6214 | -23.5527
14.6075 | -34.6218 | -23.5609
14.6136 | -34.6223 | -23.5691
14.6197 | -34.6227 | -23.5774
14.6258 | -34.6232 | -23.5858
14.6319 | -34.6236 | -23.5939
14.638 | -34.6241 | -23.6022
14.6441 | -34.6245 | -23.6104
14.6502 | -34.625 | -23.6186
14.6563 | -34.6255 | -23.6268
14.6624 | -34.6259 | -23.6349
14.6685 | -34.6264 | -23.6429
14.6746 | -34.6269 | -23.6512
14.6807 | -34.6273 | -23.6593
14.6869 | -34.6278 | -23.6674
14.693 | -34.6283 | -23.6756
14.6991 | -34.6287 | -23.6838
14.7052 | -34.6292 | -23.6918
14.7113 | -34.6297 | -23.7001
14.7174 | -34.6301 | -23.7083
14.7235 | -34.6306 | -23.7164
14.7296 | -34.6311 | -23.7243
14.7357 | -34.6315 | -23.7323
14.7418 | -34.632 | -23.7405
14.7479 | -34.6325 | -23.7486
14.754 | -34.6329 | -23.7568
14.7601 | -34.6334 | -23.7649
14.7662 | -34.6339 | -23.773
14.7723 | -34.6343 | -23.7811
14.7784 | -34.6348 | -23.7894
14.7845 | -34.6353 | -23.7976
14.7906 | -34.6357 | -23.8059
14.7967 | -34.6362 | -23.8141
14.8028 | -34.6366 | -23.8222
14.8089 | -34.6371 | -23.8302
14.815 | -34.6376 | -23.8384
14.8211 | -34.6381 | -23.8466
14.8272 | -34.6385 | -23.855
14.8334 | -34.639 | -23.8631
14.8395 | -34.6395 | -23.8713
14.8456 | -34.64 | -23.8794
14.8517 | -34.6404 | -23.8876
14.8578 | -34.6409 | -23.8959
14.8639 | -34.6413 | -23.9041
14.87 | -34.6418 | -23.9122
14.8761 | -34.6423 | -23.9203
14.8822 | -34.6428 | -23.9285
14.8883 | -34.6432 | -23.9368
14.8944 | -34.6437 | -23.9449
14.9005 | -34.6442 | -23.953
14.9066 | -34.6446 | -23.9612
14.9127 | -34.6451 | -23.9691
14.9188 | -34.6455 | -23.9772
14.9249 | -34.646 | -23.9854
14.931 | -34.6465 | -23.9934
14.9371 | -34.647 | -24.0014
14.9432 | -34.6474 | -24.0097
14.9493 | -34.6479 | -24.0178
14.9554 | -34.6483 | -24.0259
14.9615 | -34.6488 | -24.0341
14.9676 | -34.6493 | -24.0423
14.9738 | -34.6497 | -24.0503
14.9799 | -34.6502 | -24.0585
14.986 | -34.6507 | -24.0664
14.9921 | -34.6512 | -24.0745
14.9982 | -34.6516 | -24.0826
15.0043 | -34.6521 | -24.0906
15.0104 | -34.6526 | -24.0987
15.0165 | -34.653 | -24.1069
15.0226 | -34.6535 | -24.1148
15.0287 | -34.654 | -24.123
15.0348 | -34.6544 | -24.1311
15.0409 | -34.6549 | -24.1393
15.047 | -34.6554 | -24.1476
15.0531 | -34.6559 | -24.1556
15.0592 | -34.6563 | -24.1638
15.0653 | -34.6568 | -24.1721
15.0714 | -34.6572 | -24.18
15.0775 | -34.6577 | -24.188
15.0836 | -34.6582 | -24.1961
15.0897 | -34.6587 | -24.2041
15.0958 | -34.6591 | -24.2122
15.1019 | -34.6596 | -24.2203
15.108 | -34.66 | -24.2283
15.1141 | -34.6605 | -24.2364
15.1203 | -34.661 | -24.2446
15.1264 | -34.6615 | -24.2527
15.1325 | -34.662 | -24.2609
15.1386 | -34.6624 | -24.2689
15.1447 | -34.6629 | -24.277
15.1508 | -34.6634 | -24.285
15.1569 | -34.6639 | -24.2932
15.163 | -34.6643 | -24.3009
15.1691 | -34.6648 | -24.3088
15.1752 | -34.6653 | -24.3169
15.1813 | -34.6657 | -24.3249
15.1874 | -34.6662 | -24.3328
15.1935 | -34.6667 | -24.3408
15.1996 | -34.6672 | -24.3489
15.2057 | -34.6677 | -24.3568
15.2118 | -34.6681 | -24.3649
15.2179 | -34.6686 | -24.373
15.224 | -34.6691 | -24.3808
15.2301 | -34.6696 | -24.3886
15.2362 | -34.6701 | -24.3967
15.2423 | -34.6705 | -24.4048
15.2484 | -34.671 | -24.4129
15.2545 | -34.6715 | -24.4209
15.2607 | -34.672 | -24.4289
15.2668 | -34.6724 | -24.4367
15.2729 | -34.6729 | -24.4447
15.279 | -34.6734 | -24.4529
15.2851 | -34.6739 | -24.4609
15.2912 | -34.6744 | -24.4688
15.2973 | -34.6748 | -24.4771
15.3034 | -34.6753 | -24.4849
15.3095 | -34.6758 | -24.493
15.3156 | -34.6762 | -24.5008
15.3217 | -34.6767 | -24.5088
15.3278 | -34.6772 | -24.5169
15.3339 | -34.6777 | -24.5247
15.34 | -34.6782 | -24.5326
15.3461 | -34.6787 | -24.5405
15.3522 | -34.6792 | -24.5483
15.3583 | -34.6796 | -24.5564
15.3644 | -34.6801 | -24.5643
15.3705 | -34.6805 | -24.5721
15.3766 | -34.681 | -24.5801
15.3827 | -34.6815 | -24.588
15.3888 | -34.682 | -24.596
15.3949 | -34.6825 | -24.604
15.401 | -34.6829 | -24.6122
15.4072 | -34.6834 | -24.6198
15.4133 | -34.6839 | -24.6278
15.4194 | -34.6843 | -24.6358
15.4255 | -34.6848 | -24.6438
15.4316 | -34.6853 | -24.6518
15.4377 | -34.6858 | -24.6599
15.4438 | -34.6863 | -24.6679
15.4499 | -34.6867 | -24.676
15.456 | -34.6872 | -24.684
15.4621 | -34.6877 | -24.692
15.4682 | -34.6881 | -24.7001
15.4743 | -34.6886 | -24.708
15.4804 | -34.6891 | -24.7161
15.4865 | -34.6895 | -24.724
15.4926 | -34.69 | -24.732
15.4987 | -34.6905 | -24.7398
15.5048 | -34.691 | -24.7478
15.5109 | -34.6914 | -24.756
15.517 | -34.6919 | -24.764
15.5231 | -34.6924 | -24.772
15.5292 | -34.6928 | -24.7799
15.5353 | -34.6933 | -24.7879
15.5414 | -34.6938 | -24.7957
15.5476 | -34.6943 | -24.8034
15.5537 | -34.6948 | -24.8112
15.5598 | -34.6953 | -24.8192
15.5659 | -34.6957 | -24.8272
15.572 | -34.6962 | -24.8352
15.5781 | -34.6967 | -24.843
15.5842 | -34.6972 | -24.8508
15.5903 | -34.6977 | -24.8586
15.5964 | -34.6983 | -24.8665
15.6025 | -34.6987 | -24.8743
15.6086 | -34.6991 | -24.8823
15.6147 | -34.6996 | -24.8903
15.6208 | -34.7001 | -24.898
15.6269 | -34.7006 | -24.9058
15.633 | -34.701 | -24.9139
15.6391 | -34.7015 | -24.9218
15.6452 | -34.702 | -24.9298
15.6513 | -34.7025 | -24.9376
15.6574 | -34.703 | -24.9456
15.6635 | -34.7035 | -24.9535
15.6696 | -34.704 | -24.9615
15.6757 | -34.7044 | -24.9696
15.6818 | -34.7049 | -24.9775
15.688 | -34.7054 | -24.9853
15.6941 | -34.7059 | -24.993
15.7002 | -34.7063 | -25.0009
15.7063 | -34.7068 | -25.0088
15.7124 | -34.7073 | -25.0169
15.7185 | -34.7078 | -25.0248
15.7246 | -34.7083 | -25.0327
15.7307 | -34.7088 | -25.0405
15.7368 | -34.7092 | -25.0486
15.7429 | -34.7097 | -25.0567
15.749 | -34.7102 | -25.0646
15.7551 | -34.7107 | -25.0724
15.7612 | -34.7112 | -25.0802
15.7673 | -34.7117 | -25.0882
15.7734 | -34.7122 | -25.0961
15.7795 | -34.7126 | -25.1038
15.7856 | -34.7131 | -25.1116
15.7917 | -34.7136 | -25.1193
15.7978 | -34.7141 | -25.1272
15.8039 | -34.7146 | -25.1351
15.81 | -34.7151 | -25.1431
15.8161 | -34.7155 | -25.1509
15.8222 | -34.716 | -25.1587
15.8283 | -34.7165 | -25.1663
15.8345 | -34.717 | -25.1741
15.8406 | -34.7175 | -25.1819
15.8467 | -34.718 | -25.1897
15.8528 | -34.7185 | -25.1976
15.8589 | -34.719 | -25.2054
15.865 | -34.7195 | -25.2131
15.8711 | -34.72 | -25.2209
15.8772 | -34.7205 | -25.2287
15.8833 | -34.721 | -25.2366
15.8894 | -34.7214 | -25.2444
15.8955 | -34.7219 | -25.2521
15.9016 | -34.7224 | -25.2598
15.9077 | -34.7229 | -25.2676
15.9138 | -34.7233 | -25.2754
15.9199 | -34.7238 | -25.2832
15.926 | -34.7243 | -25.291
15.9321 | -34.7247 | -25.2989
15.9382 | -34.7252 | -25.3064
15.9443 | -34.7256 | -25.3142
15.9504 | -34.7261 | -25.3222
15.9565 | -34.7266 | -25.33
15.9626 | -34.7271 | -25.3378
15.9687 | -34.7276 | -25.3454
15.9749 | -34.7281 | -25.3532
15.981 | -34.7286 | -25.361
15.9871 | -34.729 | -25.3688
15.9932 | -34.7295 | -25.3765
15.9993 | -34.73 | -25.3843
16.0054 | -34.7305 | -25.3918
16.0115 | -34.7311 | -25.3997
16.0176 | -34.7316 | -25.4075
16.0237 | -34.7321 | -25.4153
16.0298 | -34.7325 | -25.423
16.0359 | -34.733 | -25.4307
16.042 | -34.7335 | -25.4384
16.0481 | -34.734 | -25.4462
16.0542 | -34.7345 | -25.4541
16.0603 | -34.7349 | -25.4618
16.0664 | -34.7354 | -25.4696
16.0725 | -34.7359 | -25.4771
16.0786 | -34.7364 | -25.4847
16.0847 | -34.7368 | -25.4926
16.0908 | -34.7373 | -25.5003
16.0969 | -34.7378 | -25.5082
16.103 | -34.7383 | -25.5159
16.1091 | -34.7388 | -25.5235
16.1152 | -34.7393 | -25.5311
16.1214 | -34.7398 | -25.5388
16.1275 | -34.7403 | -25.5466
16.1336 | -34.7407 | -25.5544
16.1397 | -34.7412 | -25.5621
16.1458 | -34.7417 | -25.5697
16.1519 | -34.7422 | -25.5776
16.158 | -34.7427 | -25.5853
16.1641 | -34.7432 | -25.5932
16.1702 | -34.7436 | -25.6009
16.1763 | -34.7441 | -25.6086
16.1824 | -34.7446 | -25.6161
16.1885 | -34.7451 | -25.6237
16.1946 | -34.7456 | -25.6315
16.2007 | -34.7461 | -25.6391
16.2068 | -34.7466 | -25.6468
16.2129 | -34.7471 | -25.6545
16.219 | -34.7476 | -25.6621
16.2251 | -34.7481 | -25.6697
16.2312 | -34.7486 | -25.6774
16.2373 | -34.7491 | -25.685
16.2434 | -34.7496 | -25.6927
16.2495 | -34.7501 | -25.7002
16.2556 | -34.7506 | -25.708
16.2618 | -34.7511 | -25.7156
16.2679 | -34.7515 | -25.7234
16.274 | -34.752 | -25.731
16.2801 | -34.7525 | -25.7386
16.2862 | -34.753 | -25.7461
16.2923 | -34.7535 | -25.7539
16.2984 | -34.754 | -25.7617
16.3045 | -34.7545 | -25.7694
16.3106 | -34.7549 | -25.7771
16.3167 | -34.7554 | -25.7847
16.3228 | -34.7559 | -25.7923
16.3289 | -34.7564 | -25.8001
16.335 | -34.7569 | -25.8078
16.3411 | -34.7574 | -25.8154
16.3472 | -34.7578 | -25.823
16.3533 | -34.7583 | -25.8307
16.3594 | -34.7588 | -25.8381
16.3655 | -34.7593 | -25.8458
16.3716 | -34.7598 | -25.8534
16.3777 | -34.7603 | -25.861
16.3838 | -34.7608 | -25.8686
16.3899 | -34.7613 | -25.876
16.396 | -34.7618 | -25.8838
16.4021 | -34.7622 | -25.8914
16.4083 | -34.7627 | -25.899
16.4144 | -34.7632 | -25.9068
16.4205 | -34.7637 | -25.9145
16.4266 | -34.7642 | -25.9219
16.4327 | -34.7647 | -25.9296
16.4388 | -34.7651 | -25.9373
16.4449 | -34.7656 | -25.9449
16.451 | -34.7661 | -25.9526
16.4571 | -34.7666 | -25.9603
16.4632 | -34.767 | -25.9677
16.4693 | -34.7675 | -25.9753
16.4754 | -34.768 | -25.9831
16.4815 | -34.7685 | -25.9907
16.4876 | -34.769 | -25.9982
16.4937 | -34.7694 | -26.0057
16.4998 | -34.7699 | -26.013
16.5059 | -34.7704 | -26.0206
16.512 | -34.7709 | -26.0284
16.5181 | -34.7714 | -26.0361
16.5242 | -34.7719 | -26.0438
16.5303 | -34.7723 | -26.0513
16.5364 | -34.7728 | -26.0588
16.5425 | -34.7733 | -26.0664
16.5487 | -34.7738 | -26.074
16.5548 | -34.7743 | -26.0816
16.5609 | -34.7748 | -26.0893
16.567 | -34.7753 | -26.0966
16.5731 | -34.7758 | -26.1043
16.5792 | -34.7762 | -26.1121
16.5853 | -34.7767 | -26.1197
16.5914 | -34.7772 | -26.1274
16.5975 | -34.7777 | -26.1351
16.6036 | -34.7781 | -26.1424
16.6097 | -34.7786 | -26.1501
16.6158 | -34.7791 | -26.1577
16.6219 | -34.7796 | -26.1653
16.628 | -34.7801 | -26.173
16.6341 | -34.7806 | -26.1805
16.6402 | -34.781 | -26.1878
16.6463 | -34.7815 | -26.1954
16.6524 | -34.782 | -26.203
16.6585 | -34.7825 | -26.2107
16.6646 | -34.783 | -26.2183
16.6707 | -34.7835 | -26.2257
16.6768 | -34.784 | -26.2332
16.6829 | -34.7845 | -26.2407
16.689 | -34.7849 | -26.2483
16.6952 | -34.7854 | -26.2558
16.7013 | -34.7859 | -26.2633
16.7074 | -34.7864 | -26.2707
16.7135 | -34.7868 | -26.2781
16.7196 | -34.7873 | -26.2857
16.7257 | -34.7878 | -26.2933
16.7318 | -34.7883 | -26.3008
16.7379 | -34.7888 | -26.3082
16.744 | -34.7893 | -26.3156
16.7501 | -34.7897 | -26.3232
16.7562 | -34.7902 | -26.3308
16.7623 | -34.7907 | -26.3383
16.7684 | -34.7912 | -26.3459
16.7745 | -34.7917 | -26.3535
16.7806 | -34.7921 | -26.3607
16.7867 | -34.7926 | -26.3682
16.7928 | -34.7931 | -26.3758
16.7989 | -34.7936 | -26.3832
16.805 | -34.7941 | -26.3909
16.8111 | -34.7946 | -26.3984
16.8172 | -34.7951 | -26.406
16.8233 | -34.7955 | -26.4136
16.8294 | -34.796 | -26.4211
16.8356 | -34.7965 | -26.4288
16.8417 | -34.797 | -26.4362
16.8478 | -34.7975 | -26.4435
16.8539 | -34.798 | -26.4508
16.86 | -34.7985 | -26.4584
16.8661 | -34.799 | -26.4659
16.8722 | -34.7994 | -26.4735
16.8783 | -34.7999 | -26.4812
16.8844 | -34.8004 | -26.4885
16.8905 | -34.8009 | -26.4961
16.8966 | -34.8014 | -26.5036
16.9027 | -34.8019 | -26.5111
16.9088 | -34.8023 | -26.5187
16.9149 | -34.8028 | -26.5263
16.921 | -34.8033 | -26.5336
16.9271 | -34.8038 | -26.5413
16.9332 | -34.8043 | -26.5487
16.9393 | -34.8047 | -26.5563
16.9454 | -34.8052 | -26.5636
16.9515 | -34.8057 | -26.5712
16.9576 | -34.8062 | -26.5785
16.9637 | -34.8067 | -26.5858
16.9698 | -34.8072 | -26.5932
16.9759 | -34.8076 | -26.6007
16.9821 | -34.8081 | -26.6082
16.9882 | -34.8086 | -26.6155
16.9943 | -34.8091 | -26.6229
17.0004 | -34.8095 | -26.6303
17.0065 | -34.81 | -26.6377
17.0126 | -34.8105 | -26.6452
17.0187 | -34.811 | -26.6528
17.0248 | -34.8115 | -26.6601
17.0309 | -34.812 | -26.6676
17.037 | -34.8125 | -26.6749
17.0431 | -34.8129 | -26.6825
17.0492 | -34.8134 | -26.69
17.0553 | -34.8139 | -26.6976
17.0614 | -34.8144 | -26.7048
17.0675 | -34.8149 | -26.7122
17.0736 | -34.8153 | -26.7196
17.0797 | -34.8158 | -26.7271
17.0858 | -34.8163 | -26.7346
17.0919 | -34.8168 | -26.742
17.098 | -34.8173 | -26.7493
17.1041 | -34.8178 | -26.7566
17.1102 | -34.8183 | -26.7641
17.1163 | -34.8187 | -26.7715
17.1225 | -34.8192 | -26.7789
17.1286 | -34.8197 | -26.7861
17.1347 | -34.8202 | -26.7936
17.1408 | -34.8207 | -26.801
17.1469 | -34.8212 | -26.8085
17.153 | -34.8217 | -26.816
17.1591 | -34.8221 | -26.8234
17.1652 | -34.8226 | -26.8306
17.1713 | -34.8231 | -26.838
17.1774 | -34.8235 | -26.8454
17.1835 | -34.824 | -26.8527
17.1896 | -34.8245 | -26.8601
17.1957 | -34.825 | -26.8675
17.2018 | -34.8255 | -26.8745
17.2079 | -34.826 | -26.882
17.214 | -34.8264 | -26.8895
17.2201 | -34.8269 | -26.897
17.2262 | -34.8274 | -26.9044
17.2323 | -34.8279 | -26.9118
17.2384 | -34.8284 | -26.9192
17.2445 | -34.8288 | -26.9264
17.2506 | -34.8293 | -26.9339
17.2567 | -34.8298 | -26.9413
17.2628 | -34.8303 | -26.9487
17.269 | -34.8308 | -26.9561
17.2751 | -34.8312 | -26.9635
17.2812 | -34.8317 | -26.9711
17.2873 | -34.8322 | -26.9785
17.2934 | -34.8327 | -26.9859
17.2995 | -34.8332 | -26.9934
17.3056 | -34.8336 | -27.0007
17.3117 | -34.8341 | -27.0082
17.3178 | -34.8346 | -27.0156
17.3239 | -34.8351 | -27.023
17.33 | -34.8355 | -27.0304
17.3361 | -34.836 | -27.0378
17.3422 | -34.8365 | -27.045
17.3483 | -34.837 | -27.0525
17.3544 | -34.8375 | -27.0598
17.3605 | -34.8379 | -27.0672
17.3666 | -34.8384 | -27.0746
17.3727 | -34.8389 | -27.0819
17.3788 | -34.8394 | -27.0891
17.3849 | -34.8399 | -27.0965
17.391 | -34.8403 | -27.1039
17.3971 | -34.8408 | -27.1112
17.4032 | -34.8413 | -27.1185
17.4094 | -34.8418 | -27.1258
17.4155 | -34.8423 | -27.1331
17.4216 | -34.8427 | -27.1404
17.4277 | -34.8432 | -27.1479
17.4338 | -34.8437 | -27.1552
17.4399 | -34.8442 | -27.1626
17.446 | -34.8447 | -27.1697
17.4521 | -34.8452 | -27.1771
17.4582 | -34.8457 | -27.1844
17.4643 | -34.8461 | -27.1917
17.4704 | -34.8466 | -27.1991
17.4765 | -34.8471 | -27.2065
17.4826 | -34.8476 | -27.2138
17.4887 | -34.8481 | -27.2211
17.4948 | -34.8486 | -27.2286
17.5009 | -34.8491 | -27.2359
17.507 | -34.8495 | -27.2433
17.5131 | -34.85 | -27.2506
17.5192 | -34.8505 | -27.2578
17.5253 | -34.851 | -27.2652
17.5314 | -34.8514 | -27.2725
17.5375 | -34.8519 | -27.2799
17.5436 | -34.8524 | -27.2873
17.5497 | -34.8528 | -27.2947
17.5559 | -34.8533 | -27.3019
17.562 | -34.8538 | -27.3093
17.5681 | -34.8543 | -27.3166
17.5742 | -34.8548 | -27.324
17.5803 | -34.8553 | -27.3313
17.5864 | -34.8558 | -27.3384
17.5925 | -34.8562 | -27.3458
17.5986 | -34.8567 | -27.3531
17.6047 | -34.8572 | -27.3606
17.6108 | -34.8577 | -27.3679
17.6169 | -34.8582 | -27.3751
17.623 | -34.8586 | -27.3823
17.6291 | -34.8591 | -27.3896
17.6352 | -34.8596 | -27.397
17.6413 | -34.8601 | -27.4043
17.6474 | -34.8606 | -27.4117
17.6535 | -34.861 | -27.4189
17.6596 | -34.8615 | -27.426
17.6657 | -34.862 | -27.4334
17.6718 | -34.8625 | -27.4407
17.6779 | -34.8629 | -27.4479
17.684 | -34.8635 | -27.4552
17.6901 | -34.8639 | -27.4626
17.6963 | -34.8644 | -27.4697
17.7024 | -34.8649 | -27.4769
17.7085 | -34.8654 | -27.4841
17.7146 | -34.8658 | -27.4913
17.7207 | -34.8663 | -27.4985
17.7268 | -34.8668 | -27.5056
17.7329 | -34.8673 | -27.513
17.739 | -34.8678 | -27.5203
17.7451 | -34.8683 | -27.5275
17.7512 | -34.8687 | -27.5347
17.7573 | -34.8692 | -27.5422
17.7634 | -34.8697 | -27.5493
17.7695 | -34.8702 | -27.5567
17.7756 | -34.8707 | -27.564
17.7817 | -34.8711 | -27.5714
17.7878 | -34.8716 | -27.5788
17.7939 | -34.8721 | -27.5862
17.8 | -34.8726 | -27.5934
17.8061 | -34.873 | -27.6007
17.8122 | -34.8735 | -27.608
17.8183 | -34.874 | -27.6153
17.8244 | -34.8745 | -27.6226
17.8305 | -34.875 | -27.6299
17.8366 | -34.8754 | -27.637
17.8428 | -34.8759 | -27.6443
17.8489 | -34.8764 | -27.6516
17.855 | -34.8769 | -27.6587
17.8611 | -34.8773 | -27.6659
17.8672 | -34.8778 | -27.673
17.8733 | -34.8783 | -27.6803
17.8794 | -34.8788 | -27.6876
17.8855 | -34.8793 | -27.6949
17.8916 | -34.8797 | -27.7022
17.8977 | -34.8802 | -27.7095
17.9038 | -34.8807 | -27.7166
17.9099 | -34.8812 | -27.7238
17.916 | -34.8816 | -27.7312
17.9221 | -34.8821 | -27.7385
17.9282 | -34.8826 | -27.7457
17.9343 | -34.8831 | -27.753
17.9404 | -34.8836 | -27.7601
17.9465 | -34.884 | -27.7674
17.9526 | -34.8845 | -27.7747
17.9587 | -34.885 | -27.782
17.9648 | -34.8855 | -27.7892
17.9709 | -34.8859 | -27.7965
17.977 | -34.8864 | -27.8036
17.9832 | -34.8869 | -27.8109
17.9893 | -34.8874 | -27.8182
17.9954 | -34.8878 | -27.8255
18.0015 | -34.8883 | -27.8328
18.0076 | -34.8888 | -27.8399
18.0137 | -34.8892 | -27.8471
18.0198 | -34.8897 | -27.8543
18.0259 | -34.8902 | -27.8616
18.032 | -34.8907 | -27.8688
18.0381 | -34.8911 | -27.8761
18.0442 | -34.8916 | -27.8833
18.0503 | -34.8921 | -27.8906
18.0564 | -34.8926 | -27.8978
18.0625 | -34.8931 | -27.9051
18.0686 | -34.8935 | -27.9125
18.0747 | -34.894 | -27.9198
18.0808 | -34.8945 | -27.927
18.0869 | -34.895 | -27.9343
18.093 | -34.8954 | -27.9416
18.0991 | -34.8959 | -27.9489
18.1052 | -34.8964 | -27.9561
18.1113 | -34.8969 | -27.9633
18.1174 | -34.8973 | -27.9704
18.1236 | -34.8978 | -27.9776
18.1297 | -34.8983 | -27.9849
18.1358 | -34.8988 | -27.9922
18.1419 | -34.8992 | -27.9994
18.148 | -34.8997 | -28.0066
18.1541 | -34.9002 | -28.0136
18.1602 | -34.9007 | -28.0209
18.1663 | -34.9012 | -28.028
18.1724 | -34.9016 | -28.0353
18.1785 | -34.9021 | -28.0425
18.1846 | -34.9026 | -28.0496
18.1907 | -34.9031 | -28.0568
18.1968 | -34.9035 | -28.064
18.2029 | -34.904 | -28.0713
18.209 | -34.9045 | -28.0787
18.2151 | -34.9049 | -28.086
18.2212 | -34.9054 | -28.0931
18.2273 | -34.9059 | -28.1004
18.2334 | -34.9064 | -28.1076
18.2395 | -34.9068 | -28.1149
18.2456 | -34.9073 | -28.1222
18.2517 | -34.9078 | -28.1294
18.2578 | -34.9082 | -28.1365
18.2639 | -34.9087 | -28.1438
18.2701 | -34.9092 | -28.151
18.2762 | -34.9097 | -28.1581
18.2823 | -34.9102 | -28.1652
18.2884 | -34.9106 | -28.1723
18.2945 | -34.9111 | -28.1793
18.3006 | -34.9116 | -28.1866
18.3067 | -34.9121 | -28.1939
18.3128 | -34.9126 | -28.2012
18.3189 | -34.913 | -28.2085
18.325 | -34.9135 | -28.2155
18.3311 | -34.914 | -28.2226
18.3372 | -34.9144 | -28.2299
18.3433 | -34.9149 | -28.2371
18.3494 | -34.9154 | -28.2443
18.3555 | -34.9159 | -28.2514
18.3616 | -34.9163 | -28.2583
18.3677 | -34.9168 | -28.2655
18.3738 | -34.9173 | -28.2726
18.3799 | -34.9177 | -28.2799
18.386 | -34.9182 | -28.287
18.3921 | -34.9187 | -28.2943
18.3982 | -34.9191 | -28.3012
18.4043 | -34.9196 | -28.3084
18.4105 | -34.9201 | -28.3156
18.4166 | -34.9205 | -28.3229
18.4227 | -34.921 | -28.3301
18.4288 | -34.9215 | -28.3373
18.4349 | -34.9219 | -28.3442
18.441 | -34.9224 | -28.3513
18.4471 | -34.9229 | -28.3584
18.4532 | -34.9233 | -28.3657
18.4593 | -34.9238 | -28.3728
18.4654 | -34.9243 | -28.3798
18.4715 | -34.9247 | -28.3871
18.4776 | -34.9252 | -28.3944
18.4837 | -34.9257 | -28.4017
18.4898 | -34.9261 | -28.4089
18.4959 | -34.9266 | -28.4161
18.502 | -34.9271 | -28.4232
18.5081 | -34.9276 | -28.4304
18.5142 | -34.928 | -28.4376
18.5203 | -34.9285 | -28.4447
18.5264 | -34.929 | -28.4517
18.5325 | -34.9295 | -28.4589
18.5386 | -34.9299 | -28.466
18.5447 | -34.9304 | -28.4731
18.5508 | -34.9309 | -28.4802
18.557 | -34.9313 | -28.4875
18.5631 | -34.9318 | -28.4946
18.5692 | -34.9323 | -28.5017
18.5753 | -34.9327 | -28.5086
18.5814 | -34.9332 | -28.5158
18.5875 | -34.9337 | -28.5229
18.5936 | -34.9341 | -28.5301
18.5997 | -34.9346 | -28.5373
18.6058 | -34.9351 | -28.5444
18.6119 | -34.9356 | -28.5515
18.618 | -34.936 | -28.5587
18.6241 | -34.9365 | -28.5658
18.6302 | -34.937 | -28.573
18.6363 | -34.9374 | -28.5802
18.6424 | -34.9379 | -28.5872
18.6485 | -34.9384 | -28.5944
18.6546 | -34.9388 | -28.6016
18.6607 | -34.9393 | -28.6089
18.6668 | -34.9398 | -28.6161
18.6729 | -34.9402 | -28.6233
18.679 | -34.9407 | -28.6302
18.6851 | -34.9412 | -28.6373
18.6912 | -34.9416 | -28.6444
18.6974 | -34.9421 | -28.6515
18.7035 | -34.9425 | -28.6586
18.7096 | -34.943 | -28.6659
18.7157 | -34.9435 | -28.6729
18.7218 | -34.9439 | -28.68
18.7279 | -34.9444 | -28.6872
18.734 | -34.9449 | -28.6944
18.7401 | -34.9454 | -28.7015
18.7462 | -34.9458 | -28.7085
18.7523 | -34.9463 | -28.7155
18.7584 | -34.9468 | -28.7227
18.7645 | -34.9472 | -28.7298
18.7706 | -34.9477 | -28.7369
18.7767 | -34.9482 | -28.7441
18.7828 | -34.9486 | -28.751
18.7889 | -34.9491 | -28.7582
18.795 | -34.9495 | -28.7653
18.8011 | -34.95 | -28.7725
18.8072 | -34.9505 | -28.7797
18.8133 | -34.9509 | -28.7867
18.8194 | -34.9514 | -28.7937
18.8255 | -34.9519 | -28.8008
18.8316 | -34.9523 | -28.8081
18.8377 | -34.9528 | -28.8152
18.8439 | -34.9533 | -28.8224
18.85 | -34.9537 | -28.8295
18.8561 | -34.9542 | -28.8366
18.8622 | -34.9546 | -28.8438
18.8683 | -34.9551 | -28.8508
18.8744 | -34.9556 | -28.858
18.8805 | -34.956 | -28.8652
18.8866 | -34.9565 | -28.8723
18.8927 | -34.957 | -28.8791
18.8988 | -34.9575 | -28.8862
18.9049 | -34.9579 | -28.8932
18.911 | -34.9584 | -28.9003
18.9171 | -34.9588 | -28.9073
18.9232 | -34.9593 | -28.9142
18.9293 | -34.9597 | -28.9213
18.9354 | -34.9602 | -28.9285
18.9415 | -34.9607 | -28.9357
18.9476 | -34.9611 | -28.9428
18.9537 | -34.9616 | -28.95
18.9598 | -34.962 | -28.9568
18.9659 | -34.9625 | -28.9638
18.972 | -34.963 | -28.9708
18.9781 | -34.9634 | -28.978
18.9843 | -34.9639 | -28.9852
18.9904 | -34.9643 | -28.9923
18.9965 | -34.9648 | -28.9992
19.0026 | -34.9653 | -29.0063
19.0087 | -34.9657 | -29.0133
19.0148 | -34.9662 | -29.0203
19.0209 | -34.9667 | -29.0273
19.027 | -34.9671 | -29.0345
19.0331 | -34.9676 | -29.0414
19.0392 | -34.9681 | -29.0485
19.0453 | -34.9685 | -29.0555
19.0514 | -34.969 | -29.0625
19.0575 | -34.9695 | -29.0698
19.0636 | -34.9699 | -29.0766
19.0697 | -34.9704 | -29.0836
19.0758 | -34.9709 | -29.0906
19.0819 | -34.9713 | -29.0977
19.088 | -34.9718 | -29.1046
19.0941 | -34.9722 | -29.1117
19.1002 | -34.9727 | -29.1186
19.1063 | -34.9731 | -29.1257
19.1124 | -34.9736 | -29.1327
19.1185 | -34.974 | -29.1397
19.1246 | -34.9745 | -29.1468
19.1308 | -34.975 | -29.154
19.1369 | -34.9755 | -29.1608
19.143 | -34.9759 | -29.1678
19.1491 | -34.9764 | -29.1748
19.1552 | -34.9768 | -29.1819
19.1613 | -34.9773 | -29.189
19.1674 | -34.9777 | -29.196
19.1735 | -34.9782 | -29.2029
19.1796 | -34.9786 | -29.2099
19.1857 | -34.9791 | -29.2169
19.1918 | -34.9796 | -29.2241
19.1979 | -34.98 | -29.2312
19.204 | -34.9805 | -29.238
19.2101 | -34.9809 | -29.245
19.2162 | -34.9814 | -29.252
19.2223 | -34.9819 | -29.2591
19.2284 | -34.9823 | -29.2662
19.2345 | -34.9828 | -29.2733
19.2406 | -34.9832 | -29.2801
19.2467 | -34.9837 | -29.2872
19.2528 | -34.9841 | -29.2944
19.2589 | -34.9846 | -29.3013
19.265 | -34.9851 | -29.3085
19.2712 | -34.9855 | -29.3155
19.2773 | -34.986 | -29.3223
19.2834 | -34.9864 | -29.3292
19.2895 | -34.9869 | -29.3362
19.2956 | -34.9874 | -29.3432
19.3017 | -34.9878 | -29.3502
19.3078 | -34.9883 | -29.3572
19.3139 | -34.9887 | -29.364
19.32 | -34.9892 | -29.371
19.3261 | -34.9896 | -29.378
19.3322 | -34.9901 | -29.3851
19.3383 | -34.9905 | -29.3921
19.3444 | -34.991 | -29.3988
19.3505 | -34.9914 | -29.4058
19.3566 | -34.9919 | -29.4127
19.3627 | -34.9923 | -29.4197
19.3688 | -34.9928 | -29.4266
19.3749 | -34.9933 | -29.4336
19.381 | -34.9937 | -29.4403
19.3871 | -34.9942 | -29.4471
19.3932 | -34.9947 | -29.4541
19.3993 | -34.9951 | -29.4612
19.4054 | -34.9956 | -29.4683
19.4115 | -34.996 | -29.4754
19.4177 | -34.9965 | -29.4822
19.4238 | -34.9969 | -29.4892
19.4299 | -34.9974 | -29.4962
19.436 | -34.9979 | -29.5033
19.4421 | -34.9983 | -29.5102
19.4482 | -34.9988 | -29.5172
19.4543 | -34.9992 | -29.5239
19.4604 | -34.9997 | -29.5309
19.4665 | -35.0001 | -29.5379
19.4726 | -35.0006 | -29.5447
19.4787 | -35.001 | -29.5517
19.4848 | -35.0015 | -29.5586
19.4909 | -35.0019 | -29.5654
19.497 | -35.0024 | -29.5724
19.5031 | -35.0028 | -29.5792
19.5092 | -35.0032 | -29.5862
19.5153 | -35.0037 | -29.5931
19.5214 | -35.0041 | -29.5998
19.5275 | -35.0046 | -29.6068
19.5336 | -35.005 | -29.6136
19.5397 | -35.0055 | -29.6205
19.5458 | -35.006 | -29.6274
19.5519 | -35.0064 | -29.6342
19.5581 | -35.0068 | -29.6408
19.5642 | -35.0073 | -29.6477
19.5703 | -35.0077 | -29.6548
19.5764 | -35.0082 | -29.6617
19.5825 | -35.0086 | -29.6686
19.5886 | -35.0091 | -29.6755
19.5947 | -35.0095 | -29.6823
19.6008 | -35.01 | -29.6892
19.6069 | -35.0104 | -29.6959
19.613 | -35.0109 | -29.7029
19.6191 | -35.0113 | -29.7096
19.6252 | -35.0118 | -29.7164
19.6313 | -35.0122 | -29.723
19.6374 | -35.0127 | -29.7299
19.6435 | -35.0131 | -29.7369
19.6496 | -35.0136 | -29.7438
19.6557 | -35.014 | -29.7507
19.6618 | -35.0145 | -29.7573
19.6679 | -35.0149 | -29.7642
19.674 | -35.0153 | -29.7712
19.6801 | -35.0158 | -29.7781
19.6862 | -35.0163 | -29.7849
19.6923 | -35.0167 | -29.7917
19.6984 | -35.0172 | -29.7984
19.7046 | -35.0176 | -29.8053
19.7107 | -35.0181 | -29.8122
19.7168 | -35.0185 | -29.8192
19.7229 | -35.019 | -29.826
19.729 | -35.0194 | -29.8329
19.7351 | -35.0199 | -29.8396
19.7412 | -35.0203 | -29.8466
19.7473 | -35.0208 | -29.8534
19.7534 | -35.0213 | -29.8604
19.7595 | -35.0217 | -29.8672
19.7656 | -35.0222 | -29.8741
19.7717 | -35.0226 | -29.8807
19.7778 | -35.023 | -29.8875
19.7839 | -35.0235 | -29.8942
19.79 | -35.024 | -29.901
19.7961 | -35.0244 | -29.908
19.8022 | -35.0248 | -29.9146
19.8083 | -35.0253 | -29.9215
19.8144 | -35.0257 | -29.9283
19.8205 | -35.0262 | -29.9351
19.8266 | -35.0266 | -29.942
19.8327 | -35.0271 | -29.9489
19.8388 | -35.0275 | -29.9558
19.845 | -35.028 | -29.9626
19.8511 | -35.0284 | -29.9694
19.8572 | -35.0288 | -29.9762
19.8633 | -35.0293 | -29.983
19.8694 | -35.0297 | -29.9899
19.8755 | -35.0302 | -29.9966
19.8816 | -35.0306 | -30.0036
19.8877 | -35.0311 | -30.0105
19.8938 | -35.0315 | -30.0173
19.8999 | -35.032 | -30.0242
19.906 | -35.0324 | -30.0312
19.9121 | -35.0329 | -30.038
19.9182 | -35.0333 | -30.0448
19.9243 | -35.0338 | -30.0515
19.9304 | -35.0342 | -30.0582
19.9365 | -35.0346 | -30.065
19.9426 | -35.0351 | -30.0716
19.9487 | -35.0355 | -30.0784
19.9548 | -35.036 | -30.0852
19.9609 | -35.0364 | -30.092
19.967 | -35.0369 | -30.0988
19.9731 | -35.0374 | -30.1056
19.9792 | -35.0378 | -30.1122
19.9853 | -35.0383 | -30.1189
19.9915 | -35.0387 | -30.1257
19.9976 | -35.0391 | -30.1325
20.0037 | -35.0396 | -30.1394
20.0098 | -35.04 | -30.1463
20.0159 | -35.0405 | -30.1532
20.022 | -35.0409 | -30.16
20.0281 | -35.0414 | -30.1667
20.0342 | -35.0418 | -30.1734
20.0403 | -35.0423 | -30.1803
20.0464 | -35.0427 | -30.1872
20.0525 | -35.0431 | -30.1937
20.0586 | -35.0436 | -30.2006
20.0647 | -35.044 | -30.2075
20.0708 | -35.0445 | -30.2143
20.0769 | -35.0449 | -30.2212
20.083 | -35.0454 | -30.2278
20.0891 | -35.0458 | -30.2346
20.0952 | -35.0463 | -30.2413
20.1013 | -35.0467 | -30.2481
20.1074 | -35.0471 | -30.2549
20.1135 | -35.0476 | -30.2616
20.1196 | -35.048 | -30.2681
20.1257 | -35.0484 | -30.2749
20.1319 | -35.0489 | -30.2817
20.138 | -35.0494 | -30.2885
20.1441 | -35.0498 | -30.2953
20.1502 | -35.0502 | -30.3021
20.1563 | -35.0507 | -30.3089
20.1624 | -35.0511 | -30.3157
20.1685 | -35.0515 | -30.3224
20.1746 | -35.052 | -30.3291
20.1807 | -35.0524 | -30.3359
20.1868 | -35.0529 | -30.3428
20.1929 | -35.0533 | -30.3493
20.199 | -35.0538 | -30.3561
20.2051 | -35.0542 | -30.3628
20.2112 | -35.0547 | -30.3696
20.2173 | -35.0551 | -30.3764
20.2234 | -35.0555 | -30.3832
20.2295 | -35.056 | -30.3897
20.2356 | -35.0564 | -30.3965
20.2417 | -35.0569 | -30.4032
20.2478 | -35.0573 | -30.41
20.2539 | -35.0578 | -30.4168
20.26 | -35.0582 | -30.4233
20.2661 | -35.0587 | -30.4301
20.2723 | -35.0591 | -30.4367
20.2784 | -35.0596 | -30.4434
20.2845 | -35.06 | -30.45
20.2906 | -35.0604 | -30.4567
20.2967 | -35.0609 | -30.4634
20.3028 | -35.0613 | -30.47
20.3089 | -35.0617 | -30.4767
20.315 | -35.0622 | -30.4833
20.3211 | -35.0626 | -30.4901
20.3272 | -35.0631 | -30.4969
20.3333 | -35.0635 | -30.5034
20.3394 | -35.064 | -30.51
20.3455 | -35.0644 | -30.5167
20.3516 | -35.0649 | -30.5234
20.3577 | -35.0653 | -30.5301
20.3638 | -35.0657 | -30.5367
20.3699 | -35.0662 | -30.5432
20.376 | -35.0667 | -30.55
20.3821 | -35.0671 | -30.5568
20.3882 | -35.0676 | -30.5635
20.3943 | -35.068 | -30.5702
20.4004 | -35.0685 | -30.5765
20.4065 | -35.0689 | -30.5831
20.4126 | -35.0693 | -30.5899
20.4188 | -35.0698 | -30.5967
20.4249 | -35.0702 | -30.6034
20.431 | -35.0706 | -30.61
20.4371 | -35.0711 | -30.6166
20.4432 | -35.0715 | -30.6232
20.4493 | -35.072 | -30.63
20.4554 | -35.0724 | -30.6365
20.4615 | -35.0729 | -30.6433
20.4676 | -35.0733 | -30.65
20.4737 | -35.0737 | -30.6565
20.4798 | -35.0742 | -30.6631
20.4859 | -35.0746 | -30.67
20.492 | -35.0751 | -30.6766
20.4981 | -35.0755 | -30.6832
20.5042 | -35.076 | -30.6898
20.5103 | -35.0764 | -30.6963
20.5164 | -35.0768 | -30.703
20.5225 | -35.0773 | -30.7096
20.5286 | -35.0777 | -30.7162
20.5347 | -35.0781 | -30.7227
20.5408 | -35.0786 | -30.7292
20.5469 | -35.079 | -30.7359
20.553 | -35.0795 | -30.7425
20.5592 | -35.0799 | -30.7491
20.5653 | -35.0804 | -30.7557
20.5714 | -35.0808 | -30.7624
20.5775 | -35.0813 | -30.7686
20.5836 | -35.0817 | -30.7754
20.5897 | -35.0821 | -30.782
20.5958 | -35.0826 | -30.7888
20.6019 | -35.083 | -30.7956
20.608 | -35.0834 | -30.8024
20.6141 | -35.0839 | -30.8089
20.6202 | -35.0843 | -30.8156
20.6263 | -35.0848 | -30.8222
20.6324 | -35.0852 | -30.8287
20.6385 | -35.0857 | -30.8352
20.6446 | -35.0862 | -30.8418
20.6507 | -35.0866 | -30.8483
20.6568 | -35.087 | -30.855
20.6629 | -35.0874 | -30.8618
20.669 | -35.0879 | -30.8685
20.6751 | -35.0883 | -30.875
20.6812 | -35.0887 | -30.8813
20.6873 | -35.0892 | -30.8881
20.6934 | -35.0896 | -30.8947
20.6995 | -35.0901 | -30.9015
20.7057 | -35.0905 | -30.9082
20.7118 | -35.0909 | -30.9148
20.7179 | -35.0914 | -30.921
20.724 | -35.0918 | -30.9277
20.7301 | -35.0922 | -30.9344
20.7362 | -35.0927 | -30.9411
20.7423 | -35.0931 | -30.9476
20.7484 | -35.0936 | -30.9542
20.7545 | -35.094 | -30.9606
20.7606 | -35.0945 | -30.9672
20.7667 | -35.0949 | -30.9738
20.7728 | -35.0954 | -30.9805
20.7789 | -35.0958 | -30.9871
20.785 | -35.0962 | -30.9936
20.7911 | -35.0967 | -30.9999
20.7972 | -35.0971 | -31.0067
20.8033 | -35.0975 | -31.0133
20.8094 | -35.098 | -31.0199
20.8155 | -35.0984 | -31.0266
20.8216 | -35.0989 | -31.0331
20.8277 | -35.0993 | -31.0399
20.8338 | -35.0997 | -31.0465
20.8399 | -35.1002 | -31.0529
20.8461 | -35.1006 | -31.0596
20.8522 | -35.101 | -31.0662
20.8583 | -35.1015 | -31.0726
20.8644 | -35.1019 | -31.0792
20.8705 | -35.1023 | -31.086
20.8766 | -35.1028 | -31.0926
20.8827 | -35.1032 | -31.0992
20.8888 | -35.1036 | -31.1057
20.8949 | -35.1041 | -31.112
20.901 | -35.1045 | -31.1186
20.9071 | -35.105 | -31.1252
20.9132 | -35.1054 | -31.1318
20.9193 | -35.1058 | -31.1386
20.9254 | -35.1063 | -31.1453
20.9315 | -35.1067 | -31.1516
20.9376 | -35.1072 | -31.1583
20.9437 | -35.1076 | -31.1648
20.9498 | -35.108 | -31.1714
20.9559 | -35.1085 | -31.1781
20.962 | -35.1089 | -31.1846
20.9681 | -35.1094 | -31.191
20.9742 | -35.1098 | -31.1977
20.9803 | -35.1102 | -31.2043
20.9864 | -35.1107 | -31.2108
20.9926 | -35.1111 | -31.2174
20.9987 | -35.1115 | -31.2237
21.0048 | -35.112 | -31.23
21.0109 | -35.1124 | -31.2366
21.017 | -35.1128 | -31.2431
21.0231 | -35.1133 | -31.2496
21.0292 | -35.1137 | -31.2563
21.0353 | -35.1141 | -31.2628
21.0414 | -35.1146 | -31.2694
21.0475 | -35.115 | -31.2759
21.0536 | -35.1155 | -31.2823
21.0597 | -35.1159 | -31.2887
21.0658 | -35.1163 | -31.2952
21.0719 | -35.1168 | -31.3015
21.078 | -35.1172 | -31.3081
21.0841 | -35.1176 | -31.3144
21.0902 | -35.118 | -31.3213
21.0963 | -35.1184 | -31.3279
21.1024 | -35.1188 | -31.3347
21.1085 | -35.1192 | -31.3412
21.1146 | -35.1196 | -31.3477
21.1207 | -35.12 | -31.3544
21.1268 | -35.1205 | -31.3611
21.133 | -35.1209 | -31.3676
21.1391 | -35.1214 | -31.3741
21.1452 | -35.1218 | -31.3808
21.1513 | -35.1222 | -31.3873
21.1574 | -35.1227 | -31.3937
21.1635 | -35.1231 | -31.4003
21.1696 | -35.1235 | -31.4068
21.1757 | -35.1239 | -31.4132
21.1818 | -35.1244 | -31.4199
21.1879 | -35.1249 | -31.4264
21.194 | -35.1253 | -31.433
21.2001 | -35.1257 | -31.4396
21.2062 | -35.1262 | -31.4462
21.2123 | -35.1266 | -31.4524
21.2184 | -35.127 | -31.4588
21.2245 | -35.1274 | -31.4654
21.2306 | -35.1279 | -31.4718
21.2367 | -35.1283 | -31.4782
21.2428 | -35.1287 | -31.4846
21.2489 | -35.1292 | -31.491
21.255 | -35.1296 | -31.4976
21.2611 | -35.13 | -31.5041
21.2672 | -35.1304 | -31.5108
21.2733 | -35.1308 | -31.5174
21.2795 | -35.1313 | -31.5237
21.2856 | -35.1317 | -31.5304
21.2917 | -35.1321 | -31.5367
21.2978 | -35.1326 | -31.543
21.3039 | -35.133 | -31.5496
21.31 | -35.1335 | -31.5562
21.3161 | -35.134 | -31.5625
21.3222 | -35.1344 | -31.5691
21.3283 | -35.1348 | -31.5756
21.3344 | -35.1353 | -31.5819
21.3405 | -35.1357 | -31.5885
21.3466 | -35.1361 | -31.5949
21.3527 | -35.1365 | -31.6012
21.3588 | -35.137 | -31.6076
21.3649 | -35.1374 | -31.614
21.371 | -35.1379 | -31.6205
21.3771 | -35.1383 | -31.6269
21.3832 | -35.1387 | -31.6333
21.3893 | -35.1392 | -31.6396
21.3954 | -35.1396 | -31.6463
21.4015 | -35.14 | -31.6527
21.4076 | -35.1405 | -31.6589
21.4137 | -35.1409 | -31.6655
21.4199 | -35.1414 | -31.6721
21.426 | -35.1418 | -31.678
21.4321 | -35.1422 | -31.6843
21.4382 | -35.1427 | -31.6908
21.4443 | -35.1431 | -31.6971
21.4504 | -35.1435 | -31.7037
21.4565 | -35.144 | -31.7099
21.4626 | -35.1444 | -31.7164
21.4687 | -35.1449 | -31.7231
21.4748 | -35.1453 | -31.7293
21.4809 | -35.1457 | -31.7356
21.487 | -35.1461 | -31.7417
21.4931 | -35.1466 | -31.748
21.4992 | -35.147 | -31.7544
21.5053 | -35.1474 | -31.7608
21.5114 | -35.1478 | -31.7674
21.5175 | -35.1482 | -31.7737
21.5236 | -35.1487 | -31.7804
21.5297 | -35.1491 | -31.7863
21.5358 | -35.1496 | -31.7927
21.5419 | -35.15 | -31.7993
21.548 | -35.1504 | -31.8059
21.5541 | -35.1508 | -31.8127
21.5602 | -35.1513 | -31.8194
21.5664 | -35.1517 | -31.8257
21.5725 | -35.1521 | -31.8322
21.5786 | -35.1525 | -31.8386
21.5847 | -35.1529 | -31.8453
21.5908 | -35.1534 | -31.8517
21.5969 | -35.1538 | -31.858
21.603 | -35.1542 | -31.8645
21.6091 | -35.1546 | -31.871
21.6152 | -35.155 | -31.8774
21.6213 | -35.1554 | -31.8839
21.6274 | -35.1559 | -31.8906
21.6335 | -35.1562 | -31.8968
21.6396 | -35.1567 | -31.9035
21.6457 | -35.1571 | -31.9099
21.6518 | -35.1575 | -31.9162
21.6579 | -35.158 | -31.9227
21.664 | -35.1584 | -31.9295
21.6701 | -35.1588 | -31.9362
21.6762 | -35.1592 | -31.9427
21.6823 | -35.1596 | -31.9493
21.6884 | -35.16 | -31.9557
21.6945 | -35.1605 | -31.9623
21.7006 | -35.1609 | -31.9688
21.7068 | -35.1614 | -31.9751
21.7129 | -35.1618 | -31.9815
21.719 | -35.1622 | -31.9879
21.7251 | -35.1626 | -31.9945
21.7312 | -35.163 | -32.0008
21.7373 | -35.1635 | -32.0073
21.7434 | -35.1639 | -32.0137
21.7495 | -35.1644 | -32.0202
21.7556 | -35.1648 | -32.0266
21.7617 | -35.1652 | -32.0332
21.7678 | -35.1657 | -32.0397
21.7739 | -35.1661 | -32.0459
21.78 | -35.1665 | -32.0522
21.7861 | -35.1669 | -32.0588
21.7922 | -35.1674 | -32.0652
21.7983 | -35.1678 | -32.0716
21.8044 | -35.1682 | -32.0783
21.8105 | -35.1686 | -32.0845
21.8166 | -35.169 | -32.0911
21.8227 | -35.1695 | -32.0974
21.8288 | -35.1699 | -32.1039
21.8349 | -35.1704 | -32.1104
21.841 | -35.1708 | -32.1168
21.8471 | -35.1712 | -32.123
21.8533 | -35.1717 | -32.1293
21.8594 | -35.1721 | -32.1356
21.8655 | -35.1725 | -32.1421
21.8716 | -35.1729 | -32.1484
21.8777 | -35.1733 | -32.1547
21.8838 | -35.1737 | -32.1609
21.8899 | -35.1741 | -32.1674
21.896 | -35.1745 | -32.1739
21.9021 | -35.175 | -32.1807
21.9082 | -35.1754 | -32.1871
21.9143 | -35.1759 | -32.1936
21.9204 | -35.1763 | -32.1999
21.9265 | -35.1767 | -32.2063
21.9326 | -35.1771 | -32.2124
21.9387 | -35.1776 | -32.2186
21.9448 | -35.178 | -32.2252
21.9509 | -35.1784 | -32.2315
21.957 | -35.1789 | -32.2379
21.9631 | -35.1793 | -32.2442
21.9692 | -35.1798 | -32.2506
21.9753 | -35.1802 | -32.2569
21.9814 | -35.1806 | -32.2634
21.9875 | -35.181 | -32.2693
21.9937 | -35.1814 | -32.2756
21.9998 | -35.1819 | -32.2821
22.0059 | -35.1823 | -32.2885
22.012 | -35.1827 | -32.2948
22.0181 | -35.1831 | -32.3011
22.0242 | -35.1836 | -32.3077
22.0303 | -35.184 | -32.3142
22.0364 | -35.1844 | -32.3205
22.0425 | -35.1849 | -32.3264
22.0486 | -35.1853 | -32.3327
22.0547 | -35.1857 | -32.339
22.0608 | -35.1861 | -32.3451
22.0669 | -35.1866 | -32.3515
22.073 | -35.187 | -32.358
22.0791 | -35.1875 | -32.3643
22.0852 | -35.1879 | -32.3707
22.0913 | -35.1884 | -32.3769
22.0974 | -35.1887 | -32.3832
22.1035 | -35.1891 | -32.3899
22.1096 | -35.1895 | -32.3964
22.1157 | -35.1899 | -32.4028
22.1218 | -35.1903 | -32.4091
22.1279 | -35.1907 | -32.4154
22.134 | -35.1911 | -32.4218
22.1402 | -35.1916 | -32.4279
22.1463 | -35.1919 | -32.4342
22.1524 | -35.1923 | -32.4407
22.1585 | -35.1927 | -32.447
22.1646 | -35.1932 | -32.4532
22.1707 | -35.1937 | -32.4598
22.1768 | -35.1941 | -32.466
22.1829 | -35.1945 | -32.4723
22.189 | -35.195 | -32.4788
22.1951 | -35.1954 | -32.4852
22.2012 | -35.1958 | -32.4915
22.2073 | -35.1962 | -32.498
22.2134 | -35.1967 | -32.5042
22.2195 | -35.1971 | -32.5105
22.2256 | -35.1975 | -32.5169
22.2317 | -35.198 | -32.5234
22.2378 | -35.1984 | -32.5299
22.2439 | -35.1988 | -32.5363
22.25 | -35.1992 | -32.5424
22.2561 | -35.1996 | -32.5488
22.2622 | -35.2 | -32.5552
22.2683 | -35.2004 | -32.5616
22.2744 | -35.2009 | -32.5677
22.2806 | -35.2013 | -32.5742
22.2867 | -35.2017 | -32.5804
22.2928 | -35.2022 | -32.5867
22.2989 | -35.2026 | -32.5931
22.305 | -35.203 | -32.5992
22.3111 | -35.2034 | -32.6056
22.3172 | -35.2039 | -32.6119
22.3233 | -35.2043 | -32.6181
22.3294 | -35.2047 | -32.6244
22.3355 | -35.2051 | -32.6308
22.3416 | -35.2056 | -32.637
22.3477 | -35.206 | -32.643
22.3538 | -35.2064 | -32.6497
22.3599 | -35.2068 | -32.6559
22.366 | -35.2072 | -32.6623
22.3721 | -35.2077 | -32.6684
22.3782 | -35.2081 | -32.6746
22.3843 | -35.2085 | -32.681
22.3904 | -35.2089 | -32.6874
22.3965 | -35.2093 | -32.6938
22.4026 | -35.2098 | -32.7001
22.4087 | -35.2102 | -32.7063
22.4148 | -35.2106 | -32.7127
22.4209 | -35.2111 | -32.7192
22.4271 | -35.2115 | -32.7254
22.4332 | -35.2119 | -32.7317
22.4393 | -35.2123 | -32.738
22.4454 | -35.2128 | -32.744
22.4515 | -35.2132 | -32.7502
22.4576 | -35.2136 | -32.7565
22.4637 | -35.214 | -32.7628
22.4698 | -35.2144 | -32.7691
22.4759 | -35.2149 | -32.7753
22.482 | -35.2153 | -32.7818
22.4881 | -35.2157 | -32.7882
22.4942 | -35.2161 | -32.7946
22.5003 | -35.2165 | -32.8008
22.5064 | -35.2169 | -32.8073
22.5125 | -35.2173 | -32.8132
22.5186 | -35.2178 | -32.8193
22.5247 | -35.2182 | -32.8256
22.5308 | -35.2186 | -32.8321
22.5369 | -35.2191 | -32.8386
22.543 | -35.2195 | -32.8448
22.5491 | -35.2199 | -32.851
22.5552 | -35.2203 | -32.8575
22.5613 | -35.2208 | -32.8638
22.5675 | -35.2212 | -32.8703
22.5736 | -35.2216 | -32.8767
22.5797 | -35.2221 | -32.8831
22.5858 | -35.2225 | -32.8892
22.5919 | -35.2229 | -32.8955
22.598 | -35.2233 | -32.9019
22.6041 | -35.2238 | -32.9082
22.6102 | -35.2242 | -32.9144
22.6163 | -35.2246 | -32.9206
22.6224 | -35.2251 | -32.927
22.6285 | -35.2255 | -32.9332
22.6346 | -35.2259 | -32.9394
22.6407 | -35.2263 | -32.9458
22.6468 | -35.2268 | -32.9523
22.6529 | -35.2272 | -32.9584
22.659 | -35.2276 | -32.9647
22.6651 | -35.228 | -32.971
22.6712 | -35.2285 | -32.9773
22.6773 | -35.2289 | -32.9837
22.6834 | -35.2293 | -32.9901
22.6895 | -35.2297 | -32.9962
22.6956 | -35.2301 | -33.0025
22.7017 | -35.2305 | -33.0087
22.7079 | -35.2309 | -33.0149
22.714 | -35.2313 | -33.0212
22.7201 | -35.2317 | -33.0276
22.7262 | -35.2321 | -33.0337
22.7323 | -35.2326 | -33.0399
22.7384 | -35.233 | -33.0461
22.7445 | -35.2334 | -33.0523
22.7506 | -35.2338 | -33.0585
22.7567 | -35.2342 | -33.0648
22.7628 | -35.2346 | -33.071
22.7689 | -35.235 | -33.0773
22.775 | -35.2355 | -33.0836
22.7811 | -35.2359 | -33.0898
22.7872 | -35.2363 | -33.096
22.7933 | -35.2368 | -33.1019
22.7994 | -35.2372 | -33.1082
22.8055 | -35.2376 | -33.1145
22.8116 | -35.238 | -33.1206
22.8177 | -35.2384 | -33.1269
22.8238 | -35.2389 | -33.1331
22.8299 | -35.2393 | -33.1392
22.836 | -35.2397 | -33.1455
22.8421 | -35.2401 | -33.1519
22.8482 | -35.2405 | -33.1582
22.8544 | -35.241 | -33.1645
22.8605 | -35.2414 | -33.1708
22.8666 | -35.2418 | -33.1769
22.8727 | -35.2422 | -33.1831
22.8788 | -35.2426 | -33.1892
22.8849 | -35.243 | -33.1954
22.891 | -35.2434 | -33.2016
22.8971 | -35.2438 | -33.2081
22.9032 | -35.2442 | -33.2141
22.9093 | -35.2446 | -33.2204
22.9154 | -35.245 | -33.2266
22.9215 | -35.2454 | -33.2328
22.9276 | -35.2458 | -33.2392
22.9337 | -35.2463 | -33.2452
22.9398 | -35.2467 | -33.2513
22.9459 | -35.2471 | -33.2575
22.952 | -35.2475 | -33.2637
22.9581 | -35.248 | -33.27
22.9642 | -35.2484 | -33.2762
22.9703 | -35.2488 | -33.2822
22.9764 | -35.2492 | -33.2885
22.9825 | -35.2496 | -33.2945
22.9886 | -35.25 | -33.3008
22.9948 | -35.2504 | -33.307
23.0009 | -35.2508 | -33.3133
23.007 | -35.2512 | -33.3192
23.0131 | -35.2517 | -33.3253
23.0192 | -35.2521 | -33.3316
23.0253 | -35.2525 | -33.3379
23.0314 | -35.2529 | -33.3441
23.0375 | -35.2533 | -33.3503
23.0436 | -35.2537 | -33.3563
23.0497 | -35.2541 | -33.3626
23.0558 | -35.2545 | -33.3688
23.0619 | -35.2549 | -33.375
23.068 | -35.2553 | -33.3812
23.0741 | -35.2557 | -33.3872
23.0802 | -35.2562 | -33.3935
23.0863 | -35.2565 | -33.3996
23.0924 | -35.257 | -33.4058
23.0985 | -35.2574 | -33.412
23.1046 | -35.2578 | -33.4181
23.1107 | -35.2582 | -33.4242
23.1168 | -35.2586 | -33.4304
23.1229 | -35.2591 | -33.4366
23.129 | -35.2595 | -33.4428
23.1351 | -35.2599 | -33.449
23.1413 | -35.2603 | -33.4552
23.1474 | -35.2607 | -33.4611
23.1535 | -35.2611 | -33.4673
23.1596 | -35.2615 | -33.4734
23.1657 | -35.2619 | -33.4796
23.1718 | -35.2623 | -33.4857
23.1779 | -35.2627 | -33.492
23.184 | -35.2631 | -33.4979
23.1901 | -35.2636 | -33.5042
23.1962 | -35.264 | -33.5103
23.2023 | -35.2644 | -33.5163
23.2084 | -35.2648 | -33.5225
23.2145 | -35.2652 | -33.5285
23.2206 | -35.2656 | -33.5347
23.2267 | -35.266 | -33.541
23.2328 | -35.2664 | -33.5471
23.2389 | -35.2668 | -33.5533
23.245 | -35.2672 | -33.5595
23.2511 | -35.2676 | -33.5655
23.2572 | -35.2681 | -33.5716
23.2633 | -35.2685 | -33.5779
23.2694 | -35.2689 | -33.5841
23.2755 | -35.2693 | -33.5902
23.2817 | -35.2697 | -33.5962
23.2878 | -35.2701 | -33.6021
23.2939 | -35.2705 | -33.6085
23.3 | -35.271 | -33.6147
23.3061 | -35.2714 | -33.6208
23.3122 | -35.2718 | -33.627
23.3183 | -35.2722 | -33.6331
23.3244 | -35.2726 | -33.6391
23.3305 | -35.273 | -33.6453
23.3366 | -35.2734 | -33.6515
23.3427 | -35.2738 | -33.6576
23.3488 | -35.2742 | -33.6637
23.3549 | -35.2746 | -33.6699
23.361 | -35.275 | -33.6759
23.3671 | -35.2755 | -33.682
23.3732 | -35.2759 | -33.6881
23.3793 | -35.2763 | -33.6942
23.3854 | -35.2767 | -33.7004
23.3915 | -35.2771 | -33.7064
23.3976 | -35.2775 | -33.7125
23.4037 | -35.278 | -33.7186
23.4098 | -35.2784 | -33.7248
23.4159 | -35.2788 | -33.731
23.422 | -35.2792 | -33.7371
23.4282 | -35.2796 | -33.743
23.4343 | -35.28 | -33.7492
23.4404 | -35.2805 | -33.7554
23.4465 | -35.2808 | -33.7616
23.4526 | -35.2813 | -33.7677
23.4587 | -35.2816 | -33.7738
23.4648 | -35.282 | -33.7797
23.4709 | -35.2824 | -33.7858
23.477 | -35.2829 | -33.7919
23.4831 | -35.2833 | -33.7979
23.4892 | -35.2837 | -33.804
23.4953 | -35.2841 | -33.8101
23.5014 | -35.2845 | -33.8161
23.5075 | -35.2849 | -33.8222
23.5136 | -35.2853 | -33.8284
23.5197 | -35.2857 | -33.8344
23.5258 | -35.2861 | -33.8404
23.5319 | -35.2865 | -33.8462
23.538 | -35.2869 | -33.8522
23.5441 | -35.2873 | -33.8586
23.5502 | -35.2877 | -33.8646
23.5563 | -35.2881 | -33.8705
23.5624 | -35.2885 | -33.8766
23.5686 | -35.2889 | -33.8824
23.5747 | -35.2893 | -33.8884
23.5808 | -35.2898 | -33.8943
23.5869 | -35.2902 | -33.9003
23.593 | -35.2906 | -33.9065
23.5991 | -35.291 | -33.9126
23.6052 | -35.2914 | -33.9185
23.6113 | -35.2918 | -33.9247
23.6174 | -35.2922 | -33.9307
23.6235 | -35.2926 | -33.9368
23.6296 | -35.293 | -33.9429
23.6357 | -35.2934 | -33.949
23.6418 | -35.2938 | -33.955
23.6479 | -35.2942 | -33.961
23.654 | -35.2946 | -33.9669
23.6601 | -35.295 | -33.9729
23.6662 | -35.2954 | -33.9789
23.6723 | -35.2958 | -33.9848
23.6784 | -35.2962 | -33.9909
23.6845 | -35.2966 | -33.9969
23.6906 | -35.297 | -34.0027
23.6967 | -35.2974 | -34.0087
23.7028 | -35.2978 | -34.0149
23.7089 | -35.2982 | -34.0207
23.7151 | -35.2986 | -34.0267
23.7212 | -35.299 | -34.0329
23.7273 | -35.2994 | -34.0388
23.7334 | -35.2999 | -34.045
23.7395 | -35.3003 | -34.0512
23.7456 | -35.3007 | -34.057
23.7517 | -35.3011 | -34.063
23.7578 | -35.3014 | -34.069
23.7639 | -35.3019 | -34.0749
23.77 | -35.3023 | -34.0809
23.7761 | -35.3027 | -34.087
23.7822 | -35.303 | -34.0929
23.7883 | -35.3035 | -34.0989
23.7944 | -35.3039 | -34.1049
23.8005 | -35.3043 | -34.111
23.8066 | -35.3047 | -34.117
23.8127 | -35.3051 | -34.1229
23.8188 | -35.3055 | -34.129
23.8249 | -35.3059 | -34.1349
23.831 | -35.3063 | -34.141
23.8371 | -35.3067 | -34.147
23.8432 | -35.3071 | -34.1531
23.8493 | -35.3075 | -34.1589
23.8555 | -35.3079 | -34.165
23.8616 | -35.3083 | -34.1708
23.8677 | -35.3087 | -34.1768
23.8738 | -35.3091 | -34.1827
23.8799 | -35.3095 | -34.1886
23.886 | -35.3099 | -34.1944
23.8921 | -35.3103 | -34.2004
23.8982 | -35.3107 | -34.2064
23.9043 | -35.3111 | -34.2124
23.9104 | -35.3115 | -34.2184
23.9165 | -35.3119 | -34.2244
23.9226 | -35.3123 | -34.2303
23.9287 | -35.3127 | -34.2363
23.9348 | -35.3131 | -34.2423
23.9409 | -35.3135 | -34.2485
23.947 | -35.3139 | -34.2545
23.9531 | -35.3143 | -34.2602
23.9592 | -35.3147 | -34.2661
23.9653 | -35.3151 | -34.2721
23.9714 | -35.3155 | -34.278
23.9775 | -35.3159 | -34.2839
23.9836 | -35.3163 | -34.29
23.9897 | -35.3167 | -34.2958
23.9958 | -35.3171 | -34.3017
24.002 | -35.3176 | -34.3077
24.0081 | -35.318 | -34.3136
24.0142 | -35.3183 | -34.3195
24.0203 | -35.3188 | -34.3254
24.0264 | -35.3192 | -34.3313
24.0325 | -35.3195 | -34.3371
24.0386 | -35.32 | -34.3432
24.0447 | -35.3203 | -34.3491
24.0508 | -35.3208 | -34.3551
24.0569 | -35.3211 | -34.3613
24.063 | -35.3216 | -34.3668
24.0691 | -35.322 | -34.3728
24.0752 | -35.3224 | -34.3789
24.0813 | -35.3228 | -34.3848
24.0874 | -35.3232 | -34.3907
24.0935 | -35.3236 | -34.3966
24.0996 | -35.324 | -34.4021
24.1057 | -35.3244 | -34.4081
24.1118 | -35.3248 | -34.4141
24.1179 | -35.3252 | -34.4201
24.124 | -35.3256 | -34.4261
24.1301 | -35.326 | -34.4318
24.1362 | -35.3264 | -34.4377
24.1424 | -35.3268 | -34.4438
24.1485 | -35.3272 | -34.4499
24.1546 | -35.3276 | -34.4558
24.1607 | -35.328 | -34.4619
24.1668 | -35.3284 | -34.4674
24.1729 | -35.3287 | -34.4734
24.179 | -35.3291 | -34.4793
24.1851 | -35.3295 | -34.4852
24.1912 | -35.3299 | -34.4911
24.1973 | -35.3303 | -34.4971
24.2034 | -35.3307 | -34.5027
24.2095 | -35.3311 | -34.5087
24.2156 | -35.3315 | -34.5147
24.2217 | -35.3319 | -34.5205
24.2278 | -35.3323 | -34.5264
24.2339 | -35.3327 | -34.5324
24.24 | -35.3331 | -34.538
24.2461 | -35.3335 | -34.544
24.2522 | -35.3339 | -34.5499
24.2583 | -35.3343 | -34.5558
24.2644 | -35.3347 | -34.5617
24.2705 | -35.3351 | -34.5673
24.2766 | -35.3355 | -34.5732
24.2827 | -35.3359 | -34.579
24.2889 | -35.3363 | -34.5848
24.295 | -35.3367 | -34.5905
24.3011 | -35.3371 | -34.5965
24.3072 | -35.3375 | -34.6021
24.3133 | -35.3379 | -34.608
24.3194 | -35.3382 | -34.6139
24.3255 | -35.3386 | -34.6198
24.3316 | -35.339 | -34.6257
24.3377 | -35.3394 | -34.6315
24.3438 | -35.3398 | -34.637
24.3499 | -35.3402 | -34.6429
24.356 | -35.3406 | -34.6486
24.3621 | -35.341 | -34.6544
24.3682 | -35.3413 | -34.6601
24.3743 | -35.3417 | -34.6662
24.3804 | -35.3421 | -34.6716
24.3865 | -35.3425 | -34.6775
24.3926 | -35.3429 | -34.6834
24.3987 | -35.3433 | -34.6893
24.4048 | -35.3437 | -34.6951
24.4109 | -35.3441 | -34.7008
24.417 | -35.3445 | -34.7067
24.4231 | -35.3448 | -34.7126
24.4293 | -35.3453 | -34.7184
24.4354 | -35.3457 | -34.7244
24.4415 | -35.346 | -34.7302
24.4476 | -35.3464 | -34.7359
24.4537 | -35.3468 | -34.7416
24.4598 | -35.3472 | -34.7475
24.4659 | -35.3476 | -34.7535
24.472 | -35.348 | -34.7594
24.4781 | -35.3484 | -34.7652
24.4842 | -35.3488 | -34.7708
24.4903 | -35.3492 | -34.7764
24.4964 | -35.3495 | -34.7822
24.5025 | -35.3499 | -34.7881
24.5086 | -35.3503 | -34.7939
24.5147 | -35.3507 | -34.7996
24.5208 | -35.3511 | -34.8052
24.5269 | -35.3515 | -34.811
24.533 | -35.3519 | -34.8168
24.5391 | -35.3523 | -34.8224
24.5452 | -35.3527 | -34.8281
24.5513 | -35.353 | -34.8337
24.5574 | -35.3534 | -34.8393
24.5635 | -35.3538 | -34.8451
24.5696 | -35.3542 | -34.8509
24.5758 | -35.3546 | -34.8565
24.5819 | -35.355 | -34.8622
24.588 | -35.3554 | -34.8678
24.5941 | -35.3558 | -34.8736
24.6002 | -35.3562 | -34.8794
24.6063 | -35.3566 | -34.8851
24.6124 | -35.357 | -34.8909
24.6185 | -35.3574 | -34.8968
24.6246 | -35.3577 | -34.9021
24.6307 | -35.3581 | -34.9079
24.6368 | -35.3585 | -34.9137
24.6429 | -35.3589 | -34.9193
24.649 | -35.3593 | -34.925
24.6551 | -35.3597 | -34.9308
24.6612 | -35.3601 | -34.9361
24.6673 | -35.3605 | -34.9419
24.6734 | -35.3609 | -34.9476
24.6795 | -35.3613 | -34.9535
24.6856 | -35.3617 | -34.9592
24.6917 | -35.3621 | -34.965
24.6978 | -35.3624 | -34.9704
24.7039 | -35.3628 | -34.9762
24.71 | -35.3632 | -34.9819
24.7162 | -35.3636 | -34.9877
24.7223 | -35.364 | -34.9934
24.7284 | -35.3644 | -34.9988
24.7345 | -35.3647 | -35.0044
24.7406 | -35.3652 | -35.01
24.7467 | -35.3655 | -35.0159
24.7528 | -35.3659 | -35.0216
24.7589 | -35.3663 | -35.0272
24.765 | -35.3667 | -35.0325
24.7711 | -35.3671 | -35.0382
24.7772 | -35.3675 | -35.044
24.7833 | -35.3678 | -35.0496
24.7894 | -35.3682 | -35.0554
24.7955 | -35.3686 | -35.061
24.8016 | -35.369 | -35.0665
24.8077 | -35.3694 | -35.0723
24.8138 | -35.3697 | -35.078
24.8199 | -35.3701 | -35.0836
24.826 | -35.3705 | -35.0892
24.8321 | -35.3709 | -35.095
24.8382 | -35.3713 | -35.1005
24.8443 | -35.3716 | -35.106
24.8504 | -35.372 | -35.1117
24.8565 | -35.3724 | -35.1174
24.8627 | -35.3728 | -35.1231
24.8688 | -35.3732 | -35.1286
24.8749 | -35.3736 | -35.1343
24.881 | -35.374 | -35.1401
24.8871 | -35.3744 | -35.1458
24.8932 | -35.3748 | -35.1515
24.8993 | -35.3752 | -35.157
24.9054 | -35.3755 | -35.1626
24.9115 | -35.3759 | -35.1683
24.9176 | -35.3763 | -35.174
24.9237 | -35.3767 | -35.1796
24.9298 | -35.3771 | -35.1853
24.9359 | -35.3775 | -35.191
24.942 | -35.3779 | -35.1963
24.9481 | -35.3783 | -35.202
24.9542 | -35.3786 | -35.2076
24.9603 | -35.379 | -35.2134
24.9664 | -35.3794 | -35.219
24.9725 | -35.3798 | -35.2246
24.9786 | -35.3802 | -35.23
24.9847 | -35.3806 | -35.2357
24.9908 | -35.381 | -35.2414
24.9969 | -35.3813 | -35.247
25.0031 | -35.3817 | -35.2528
25.0092 | -35.3821 | -35.2583
25.0153 | -35.3824 | -35.2639
25.0214 | -35.3828 | -35.2696
25.0275 | -35.3832 | -35.2752
25.0336 | -35.3836 | -35.2808
25.0397 | -35.384 | -35.2866
25.0458 | -35.3844 | -35.2918
25.0519 | -35.3848 | -35.2974
25.058 | -35.3851 | -35.303
25.0641 | -35.3855 | -35.3087
25.0702 | -35.3859 | -35.3142
25.0763 | -35.3862 | -35.3199
25.0824 | -35.3866 | -35.3252
25.0885 | -35.387 | -35.3309
25.0946 | -35.3873 | -35.3365
25.1007 | -35.3877 | -35.3423
25.1068 | -35.3881 | -35.3478
25.1129 | -35.3885 | -35.3536
25.119 | -35.3889 | -35.3587
25.1251 | -35.3892 | -35.3643
25.1312 | -35.3896 | -35.3698
25.1373 | -35.39 | -35.3754
25.1435 | -35.3904 | -35.3811
25.1496 | -35.3908 | -35.3866
25.1557 | -35.3911 | -35.3922
25.1618 | -35.3915 | -35.3979
25.1679 | -35.3919 | -35.4035
25.174 | -35.3923 | -35.409
25.1801 | -35.3926 | -35.4147
25.1862 | -35.393 | -35.4202
25.1923 | -35.3934 | -35.4256
25.1984 | -35.3938 | -35.4312
25.2045 | -35.3941 | -35.4368
25.2106 | -35.3945 | -35.4424
25.2167 | -35.3949 | -35.4481
25.2228 | -35.3952 | -35.4535
25.2289 | -35.3956 | -35.459
25.235 | -35.396 | -35.4646
25.2411 | -35.3964 | -35.4701
25.2472 | -35.3967 | -35.4755
25.2533 | -35.3971 | -35.481
25.2594 | -35.3975 | -35.4863
25.2655 | -35.3978 | -35.4917
25.2716 | -35.3982 | -35.4971
25.2777 | -35.3986 | -35.5026
25.2838 | -35.399 | -35.5082
25.29 | -35.3993 | -35.5137
25.2961 | -35.3997 | -35.5188
25.3022 | -35.4001 | -35.5245
25.3083 | -35.4005 | -35.5301
25.3144 | -35.4008 | -35.5356
25.3205 | -35.4012 | -35.5412
25.3266 | -35.4016 | -35.5465
25.3327 | -35.402 | -35.5521
25.3388 | -35.4023 | -35.5576
25.3449 | -35.4027 | -35.5632
25.351 | -35.4031 | -35.5686
25.3571 | -35.4035 | -35.5741
25.3632 | -35.4038 | -35.5795
25.3693 | -35.4042 | -35.5851
25.3754 | -35.4046 | -35.5906
25.3815 | -35.405 | -35.5962
25.3876 | -35.4053 | -35.6016
25.3937 | -35.4057 | -35.6071
25.3998 | -35.4061 | -35.6122
25.4059 | -35.4064 | -35.6177
25.412 | -35.4068 | -35.6232
25.4181 | -35.4072 | -35.6286
25.4242 | -35.4076 | -35.6341
25.4304 | -35.4079 | -35.6396
25.4365 | -35.4083 | -35.6446
25.4426 | -35.4087 | -35.6502
25.4487 | -35.409 | -35.6558
25.4548 | -35.4094 | -35.6613
25.4609 | -35.4097 | -35.6667
25.467 | -35.4101 | -35.672
25.4731 | -35.4105 | -35.6777
25.4792 | -35.4109 | -35.6833
25.4853 | -35.4112 | -35.6889
25.4914 | -35.4116 | -35.6943
25.4975 | -35.412 | -35.6997
25.5036 | -35.4123 | -35.7051
25.5097 | -35.4127 | -35.7104
25.5158 | -35.4131 | -35.7159
25.5219 | -35.4134 | -35.7215
25.528 | -35.4138 | -35.7268
25.5341 | -35.4142 | -35.7324
25.5402 | -35.4146 | -35.7378
25.5463 | -35.4149 | -35.7432
25.5524 | -35.4153 | -35.7487
25.5585 | -35.4157 | -35.7541
25.5646 | -35.4161 | -35.7597
25.5707 | -35.4164 | -35.7652
25.5769 | -35.4168 | -35.7705
25.583 | -35.4172 | -35.776
25.5891 | -35.4176 | -35.7816
25.5952 | -35.4179 | -35.787
25.6013 | -35.4183 | -35.7924
25.6074 | -35.4187 | -35.7976
25.6135 | -35.419 | -35.8031
25.6196 | -35.4194 | -35.8085
25.6257 | -35.4198 | -35.8139
25.6318 | -35.4202 | -35.8192
25.6379 | -35.4205 | -35.8248
25.644 | -35.4209 | -35.83
25.6501 | -35.4213 | -35.8354
25.6562 | -35.4216 | -35.8408
25.6623 | -35.422 | -35.8462
25.6684 | -35.4224 | -35.8517
25.6745 | -35.4227 | -35.8572
25.6806 | -35.4231 | -35.8623
25.6867 | -35.4235 | -35.8677
25.6928 | -35.4239 | -35.8731
25.6989 | -35.4242 | -35.8786
25.705 | -35.4246 | -35.884
25.7111 | -35.4249 | -35.8895
25.7173 | -35.4253 | -35.8946
25.7234 | -35.4257 | -35.9001
25.7295 | -35.426 | -35.9056
25.7356 | -35.4264 | -35.9111
25.7417 | -35.4268 | -35.9164
25.7478 | -35.4271 | -35.9217
25.7539 | -35.4275 | -35.927
25.76 | -35.4278 | -35.9325
25.7661 | -35.4282 | -35.9378
25.7722 | -35.4285 | -35.9431
25.7783 | -35.4289 | -35.9486
25.7844 | -35.4293 | -35.9538
25.7905 | -35.4296 | -35.9593
25.7966 | -35.43 | -35.9646
25.8027 | -35.4303 | -35.9699
25.8088 | -35.4307 | -35.9754
25.8149 | -35.4311 | -35.9809
25.821 | -35.4314 | -35.986
25.8271 | -35.4318 | -35.9914
25.8332 | -35.4321 | -35.9969
25.8393 | -35.4325 | -36.0023
25.8454 | -35.4329 | -36.0078
25.8515 | -35.4332 | -36.0132
25.8576 | -35.4336 | -36.0184
25.8638 | -35.4339 | -36.0238
25.8699 | -35.4343 | -36.029
25.876 | -35.4346 | -36.0345
25.8821 | -35.435 | -36.0398
25.8882 | -35.4353 | -36.0447
25.8943 | -35.4357 | -36.0502
25.9004 | -35.436 | -36.0555
25.9065 | -35.4364 | -36.0609
25.9126 | -35.4368 | -36.0662
25.9187 | -35.4371 | -36.0713
25.9248 | -35.4375 | -36.0765
25.9309 | -35.4378 | -36.0819
25.937 | -35.4382 | -36.0872
25.9431 | -35.4386 | -36.0926
25.9492 | -35.4389 | -36.0979
25.9553 | -35.4393 | -36.1034
25.9614 | -35.4396 | -36.1085
25.9675 | -35.44 | -36.1139
25.9736 | -35.4403 | -36.1193
25.9797 | -35.4407 | -36.1248
25.9858 | -35.441 | -36.1301
25.9919 | -35.4414 | -36.1355
25.998 | -35.4417 | -36.1407
26.0042 | -35.4421 | -36.146
26.0103 | -35.4424 | -36.1513
26.0164 | -35.4428 | -36.1567
26.0225 | -35.4431 | -36.1619
26.0286 | -35.4435 | -36.1673
26.0347 | -35.4438 | -36.1723
26.0408 | -35.4442 | -36.1777
26.0469 | -35.4446 | -36.1831
26.053 | -35.4449 | -36.1884
26.0591 | -35.4453 | -36.1937
26.0652 | -35.4456 | -36.1988
26.0713 | -35.446 | -36.2041
26.0774 | -35.4463 | -36.2094
26.0835 | -35.4467 | -36.2149
26.0896 | -35.447 | -36.2203
26.0957 | -35.4474 | -36.2256
26.1018 | -35.4477 | -36.2307
26.1079 | -35.4481 | -36.236
26.114 | -35.4485 | -36.2413
26.1201 | -35.4488 | -36.2466
26.1262 | -35.4492 | -36.2518
26.1323 | -35.4495 | -36.2572
26.1384 | -35.4498 | -36.2623
26.1445 | -35.4502 | -36.2677
26.1507 | -35.4506 | -36.2731
26.1568 | -35.4509 | -36.2785
26.1629 | -35.4513 | -36.2838
26.169 | -35.4516 | -36.289
26.1751 | -35.4519 | -36.294
26.1812 | -35.4523 | -36.2994
26.1873 | -35.4526 | -36.3048
26.1934 | -35.453 | -36.3101
26.1995 | -35.4534 | -36.3154
26.2056 | -35.4537 | -36.3205
26.2117 | -35.454 | -36.3258
26.2178 | -35.4544 | -36.3309
26.2239 | -35.4547 | -36.3362
26.23 | -35.4551 | -36.3416
26.2361 | -35.4554 | -36.3469
26.2422 | -35.4557 | -36.3519
26.2483 | -35.4561 | -36.3571
26.2544 | -35.4564 | -36.3624
26.2605 | -35.4568 | -36.3676
26.2666 | -35.4571 | -36.3729
26.2727 | -35.4575 | -36.3781
26.2788 | -35.4578 | -36.3832
26.2849 | -35.4581 | -36.3884
26.2911 | -35.4585 | -36.3936
26.2972 | -35.4588 | -36.3988
26.3033 | -35.4592 | -36.4042
26.3094 | -35.4595 | -36.4095
26.3155 | -35.4599 | -36.4146
26.3216 | -35.4602 | -36.42
26.3277 | -35.4605 | -36.4254
26.3338 | -35.4609 | -36.4306
26.3399 | -35.4612 | -36.4357
26.346 | -35.4615 | -36.4407
26.3521 | -35.4619 | -36.4461
26.3582 | -35.4623 | -36.4513
26.3643 | -35.4626 | -36.4566
26.3704 | -35.463 | -36.4619
26.3765 | -35.4633 | -36.4672
26.3826 | -35.4637 | -36.4722
26.3887 | -35.464 | -36.4776
26.3948 | -35.4644 | -36.4828
26.4009 | -35.4647 | -36.4881
26.407 | -35.465 | -36.4934
26.4131 | -35.4654 | -36.4988
26.4192 | -35.4657 | -36.5037
26.4253 | -35.4661 | -36.5088
26.4314 | -35.4664 | -36.5139
26.4376 | -35.4667 | -36.5192
26.4437 | -35.4671 | -36.5243
26.4498 | -35.4674 | -36.5296
26.4559 | -35.4677 | -36.5346
26.462 | -35.4681 | -36.5399
26.4681 | -35.4684 | -36.5451
26.4742 | -35.4688 | -36.5503
26.4803 | -35.4691 | -36.5555
26.4864 | -35.4694 | -36.5606
26.4925 | -35.4698 | -36.5658
26.4986 | -35.4701 | -36.571
26.5047 | -35.4705 | -36.5762
26.5108 | -35.4708 | -36.5813
26.5169 | -35.4712 | -36.5866
26.523 | -35.4715 | -36.5916
26.5291 | -35.4718 | -36.5969
26.5352 | -35.4722 | -36.6022
26.5413 | -35.4725 | -36.6075
26.5474 | -35.4728 | -36.6128
26.5535 | -35.4732 | -36.6182
26.5596 | -35.4735 | -36.6232
26.5657 | -35.4738 | -36.6284
26.5718 | -35.4742 | -36.6335
26.578 | -35.4745 | -36.6388
26.5841 | -35.4748 | -36.6439
26.5902 | -35.4752 | -36.6492
26.5963 | -35.4755 | -36.6542
26.6024 | -35.4758 | -36.6595
26.6085 | -35.4761 | -36.6646
26.6146 | -35.4765 | -36.67
26.6207 | -35.4768 | -36.6753
26.6268 | -35.4771 | -36.6805
26.6329 | -35.4775 | -36.6855
26.639 | -35.4778 | -36.6906
26.6451 | -35.4781 | -36.6959
26.6512 | -35.4785 | -36.701
26.6573 | -35.4788 | -36.7061
26.6634 | -35.4791 | -36.711
26.6695 | -35.4794 | -36.7163
26.6756 | -35.4798 | -36.7216
26.6817 | -35.4801 | -36.7269
26.6878 | -35.4805 | -36.7321
26.6939 | -35.4808 | -36.7374
26.7 | -35.4811 | -36.7424
26.7061 | -35.4815 | -36.7476
26.7122 | -35.4818 | -36.7528
26.7183 | -35.4821 | -36.758
26.7245 | -35.4824 | -36.7632
26.7306 | -35.4828 | -36.7683
26.7367 | -35.4831 | -36.7733
26.7428 | -35.4834 | -36.7785
26.7489 | -35.4838 | -36.7837
26.755 | -35.4841 | -36.7886
26.7611 | -35.4845 | -36.7937
26.7672 | -35.4848 | -36.7989
26.7733 | -35.4851 | -36.8041
26.7794 | -35.4855 | -36.8093
26.7855 | -35.4858 | -36.8145
26.7916 | -35.4861 | -36.8196
26.7977 | -35.4865 | -36.8248
26.8038 | -35.4868 | -36.8299
26.8099 | -35.4871 | -36.8351
26.816 | -35.4874 | -36.8403
26.8221 | -35.4877 | -36.8454
26.8282 | -35.4881 | -36.8505
26.8343 | -35.4884 | -36.8558
26.8404 | -35.4887 | -36.861
26.8465 | -35.489 | -36.8661
26.8526 | -35.4893 | -36.8711
26.8587 | -35.4896 | -36.8763
26.8649 | -35.49 | -36.8815
26.871 | -35.4903 | -36.8868
26.8771 | -35.4906 | -36.8918
26.8832 | -35.4909 | -36.8969
26.8893 | -35.4913 | -36.9021
26.8954 | -35.4916 | -36.9073
26.9015 | -35.4919 | -36.9125
26.9076 | -35.4922 | -36.9177
26.9137 | -35.4926 | -36.9227
26.9198 | -35.4929 | -36.928
26.9259 | -35.4932 | -36.933
26.932 | -35.4935 | -36.9382
26.9381 | -35.4939 | -36.9433
26.9442 | -35.4942 | -36.9484
26.9503 | -35.4945 | -36.9535
26.9564 | -35.4948 | -36.9587
26.9625 | -35.4952 | -36.9638
26.9686 | -35.4955 | -36.9689
26.9747 | -35.4958 | -36.974
26.9808 | -35.4961 | -36.979
26.9869 | -35.4964 | -36.9842
26.993 | -35.4968 | -36.9893
26.9991 | -35.4971 | -36.9944
27.0052 | -35.4974 | -36.9996
27.0114 | -35.4977 | -37.0048
27.0175 | -35.498 | -37.0097
27.0236 | -35.4983 | -37.015
27.0297 | -35.4987 | -37.0201
27.0358 | -35.499 | -37.0253
27.0419 | -35.4993 | -37.0305
27.048 | -35.4996 | -37.0357
27.0541 | -35.4999 | -37.0405
27.0602 | -35.5003 | -37.0456
27.0663 | -35.5006 | -37.0509
27.0724 | -35.5009 | -37.0558
27.0785 | -35.5013 | -37.0611
27.0846 | -35.5016 | -37.066
27.0907 | -35.5019 | -37.0712
27.0968 | -35.5022 | -37.0764
27.1029 | -35.5025 | -37.0814
27.109 | -35.5028 | -37.0866
27.1151 | -35.5032 | -37.0919
27.1212 | -35.5034 | -37.097
27.1273 | -35.5038 | -37.1021
27.1334 | -35.5041 | -37.1073
27.1395 | -35.5044 | -37.1124
27.1456 | -35.5047 | -37.1175
27.1518 | -35.5051 | -37.1227
27.1579 | -35.5054 | -37.1277
27.164 | -35.5057 | -37.1327
27.1701 | -35.506 | -37.1378
27.1762 | -35.5063 | -37.1428
27.1823 | -35.5066 | -37.1479
27.1884 | -35.5069 | -37.1531
27.1945 | -35.5072 | -37.1581
27.2006 | -35.5075 | -37.1632
27.2067 | -35.5079 | -37.1684
27.2128 | -35.5082 | -37.1735
27.2189 | -35.5085 | -37.1786
27.225 | -35.5088 | -37.1838
27.2311 | -35.5092 | -37.1886
27.2372 | -35.5095 | -37.1938
27.2433 | -35.5098 | -37.1988
27.2494 | -35.5101 | -37.204
27.2555 | -35.5105 | -37.209
27.2616 | -35.5107 | -37.214
27.2677 | -35.5111 | -37.2192
27.2738 | -35.5114 | -37.2242
27.2799 | -35.5117 | -37.2294
27.286 | -35.512 | -37.2345
27.2921 | -35.5123 | -37.2396
27.2983 | -35.5126 | -37.2445
27.3044 | -35.5129 | -37.2496
27.3105 | -35.5132 | -37.2546
27.3166 | -35.5135 | -37.2597
27.3227 | -35.5138 | -37.2648
27.3288 | -35.5142 | -37.2699
27.3349 | -35.5144 | -37.2749
27.341 | -35.5147 | -37.2799
27.3471 | -35.5151 | -37.2849
27.3532 | -35.5154 | -37.2899
27.3593 | -35.5157 | -37.2949
27.3654 | -35.516 | -37.3
27.3715 | -35.5163 | -37.305
27.3776 | -35.5166 | -37.3103
27.3837 | -35.517 | -37.3154
27.3898 | -35.5173 | -37.3204
27.3959 | -35.5176 | -37.3255
27.402 | -35.5179 | -37.3303
27.4081 | -35.5182 | -37.3355
27.4142 | -35.5185 | -37.3405
27.4203 | -35.5188 | -37.3457
27.4264 | -35.5191 | -37.3507
27.4325 | -35.5195 | -37.3558
27.4387 | -35.5197 | -37.3609
27.4448 | -35.52 | -37.366
27.4509 | -35.5203 | -37.3712
27.457 | -35.5207 | -37.3762
27.4631 | -35.521 | -37.3814
27.4692 | -35.5213 | -37.3864
27.4753 | -35.5215 | -37.3913
27.4814 | -35.5218 | -37.3964
27.4875 | -35.5221 | -37.4016
27.4936 | -35.5224 | -37.4067
27.4997 | -35.5227 | -37.4119
27.5058 | -35.5231 | -37.4171
27.5119 | -35.5234 | -37.422
27.518 | -35.5237 | -37.4271
27.5241 | -35.524 | -37.4322
27.5302 | -35.5243 | -37.4374
27.5363 | -35.5246 | -37.4424
27.5424 | -35.5249 | -37.4474
27.5485 | -35.5252 | -37.4524
27.5546 | -35.5255 | -37.4574
27.5607 | -35.5258 | -37.4624
27.5668 | -35.5261 | -37.4675
27.5729 | -35.5264 | -37.4725
27.5791 | -35.5267 | -37.4774
27.5852 | -35.527 | -37.4825
27.5913 | -35.5273 | -37.4873
27.5974 | -35.5276 | -37.4924
27.6035 | -35.5279 | -37.4975
27.6096 | -35.5282 | -37.5025
27.6157 | -35.5285 | -37.5073
27.6218 | -35.5288 | -37.5123
27.6279 | -35.5291 | -37.5172
27.634 | -35.5294 | -37.5223
27.6401 | -35.5297 | -37.5274
27.6462 | -35.53 | -37.5325
27.6523 | -35.5303 | -37.5375
27.6584 | -35.5306 | -37.5427
27.6645 | -35.5309 | -37.5478
27.6706 | -35.5313 | -37.5528
27.6767 | -35.5316 | -37.5579
27.6828 | -35.5319 | -37.563
27.6889 | -35.5322 | -37.568
27.695 | -35.5325 | -37.573
27.7011 | -35.5328 | -37.578
27.7072 | -35.5331 | -37.5831
27.7133 | -35.5334 | -37.5881
27.7194 | -35.5337 | -37.593
27.7256 | -35.534 | -37.598
27.7317 | -35.5343 | -37.603
27.7378 | -35.5346 | -37.608
27.7439 | -35.5349 | -37.6132
27.75 | -35.5352 | -37.6183
27.7561 | -35.5355 | -37.6232
27.7622 | -35.5358 | -37.6283
27.7683 | -35.5361 | -37.6335
27.7744 | -35.5364 | -37.6386
27.7805 | -35.5367 | -37.6437
27.7866 | -35.537 | -37.6488
27.7927 | -35.5373 | -37.6537
27.7988 | -35.5376 | -37.6588
27.8049 | -35.5379 | -37.6637
27.811 | -35.5382 | -37.6689
27.8171 | -35.5385 | -37.6739
27.8232 | -35.5388 | -37.6788
27.8293 | -35.5391 | -37.6838
27.8354 | -35.5394 | -37.6889
27.8415 | -35.5397 | -37.6938
27.8476 | -35.54 | -37.6988
27.8537 | -35.5403 | -37.704
27.8598 | -35.5406 | -37.7089
27.866 | -35.5409 | -37.714
27.8721 | -35.5411 | -37.719
27.8782 | -35.5414 | -37.7242
27.8843 | -35.5417 | -37.7293
27.8904 | -35.542 | -37.7343
27.8965 | -35.5423 | -37.7393
27.9026 | -35.5426 | -37.7442
27.9087 | -35.5429 | -37.7493
27.9148 | -35.5432 | -37.7542
27.9209 | -35.5435 | -37.7594
27.927 | -35.5438 | -37.7643
27.9331 | -35.5441 | -37.7694
27.9392 | -35.5444 | -37.7743
27.9453 | -35.5447 | -37.7793
27.9514 | -35.545 | -37.7843
27.9575 | -35.5453 | -37.7895
27.9636 | -35.5455 | -37.7945
27.9697 | -35.5459 | -37.7995
27.9758 | -35.5461 | -37.8045
27.9819 | -35.5464 | -37.8096
27.988 | -35.5467 | -37.8145
27.9941 | -35.547 | -37.8195
28.0002 | -35.5473 | -37.8244
28.0063 | -35.5475 | -37.8295
28.0125 | -35.5478 | -37.8343
28.0186 | -35.5481 | -37.8393
28.0247 | -35.5484 | -37.8444
28.0308 | -35.5487 | -37.8494
28.0369 | -35.549 | -37.8543
28.043 | -35.5493 | -37.8592
28.0491 | -35.5496 | -37.8642
28.0552 | -35.5498 | -37.8694
28.0613 | -35.5501 | -37.8743
28.0674 | -35.5504 | -37.8792
28.0735 | -35.5507 | -37.8842

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5563 | -43.367
21.4908 | -35.6099 | -35.3769
16.4948 | -34.9844 | -27.9854
12.6602 | -34.6185 | -21.7047
9.7171 | -34.3896 | -16.7969
7.4582 | -34.232 | -12.8101
5.7244 | -34.1208 | -9.5824
4.3936 | -34.0499 | -6.9808
3.3722 | -34.021 | -4.927
2.5883 | -34.0205 | -3.3694
1.9866 | -34.0022 | -2.1245
1.5248 | -34.0018 | -1.1268
1.1703 | -34.0058 | -0.47353
1 | -34.0067 | -0.15023
0.89822 | -33.9958 | 0.12541
0.68941 | -33.9725 | 0.70644
0.52919 | -33.9654 | 1.1615
0.40612 | -33.9664 | 1.7693
0.31171 | -33.9527 | 2.2288
0.23928 | -33.9544 | 2.9544
0.18358 | -33.9581 | 3.3747
0.14091 | -33.9593 | 4.2558
0.10819 | -33.9776 | 5.4135
0.083017 | -33.9723 | 6.9066
0.063705 | -34.0316 | 8.777
0.048923 | -34.1026 | 11.2762
0.037479 | -34.2094 | 14.4758
0.028801 | -34.3898 | 18.5091
0.022078 | -34.6933 | 23.8526
0.017023 | -35.1689 | 30.2584
0.013018 | -35.9741 | 37.9401
0.010014 | -37.1067 | 45.8433

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003673 | -10.1217
0.007347 | -26.8678
0.01102 | -35.4954
0.014693 | -39.7126
0.018366 | -42.6521
0.02204 | -45.4215
0.025713 | -46.3122
0.029386 | -47.5321
0.033059 | -48.5039
0.036733 | -49.6019
0.040406 | -50.8353
0.044079 | -51.4022
0.047752 | -53.2933
0.051426 | -54.8617
0.055099 | -55.5563
0.058772 | -55.5083
0.062445 | -56.1528
0.066119 | -56.9612
0.069792 | -57.1703
0.073465 | -57.2585
0.077139 | -59.0225
0.080812 | -60.7819
0.084485 | -60.7415
0.088158 | -61.7467
0.091832 | -62.1424
0.095505 | -61.7551
0.099178 | -61.7727
0.10285 | -61.2786
0.10652 | -61.6469
0.1102 | -63.007
0.11387 | -64.3742
0.11754 | -64.9906
0.12122 | -64.3484
0.12489 | -64.0985
0.12856 | -64.2851
0.13224 | -64.6926
0.13591 | -64.7665
0.13958 | -64.6992
0.14326 | -64.6891
0.14693 | -64.905
0.1506 | -65.7632
0.15428 | -67.0681
0.15795 | -67.275
0.16162 | -67.3818
0.1653 | -68.2442
0.16897 | -67.487
0.17264 | -67.1048
0.17632 | -68.0213
0.17999 | -69.1941
0.18366 | -69.4153
0.18734 | -69.5035
0.19101 | -69.0299
0.19468 | -69.0174
0.19836 | -69.1573
0.20203 | -68.9889
0.2057 | -70.0138
0.20938 | -70.1319
0.21305 | -70.0912
0.21672 | -70.6196
0.2204 | -70.1532
0.22407 | -70.2647
0.22774 | -70.8026
0.23142 | -70.0441
0.23509 | -69.7288
0.23876 | -70.4632
0.24244 | -71.1176
0.24611 | -72.6395
0.24978 | -73.0466
0.25345 | -72.4944
0.25713 | -72.4667
0.2608 | -72.0696
0.26448 | -72.3229
0.26815 | -72.8808
0.27182 | -73.0328
0.27549 | -72.8419
0.27917 | -73.3865
0.28284 | -73.5207
0.28652 | -72.8481
0.29019 | -73.0676
0.29386 | -73.4289
0.29753 | -72.2471
0.30121 | -73.1111
0.30488 | -73.3409
0.30855 | -73.9638
0.31223 | -74.5974
0.3159 | -73.7414
0.31957 | -74.3792
0.32325 | -74.64
0.32692 | -73.9092
0.33059 | -73.6909
0.33427 | -74.2386
0.33794 | -75.1008
0.34161 | -75.6317
0.34529 | -76.0851
0.34896 | -76.6101
0.35263 | -76.3153
0.35631 | -75.7761
0.35998 | -75.9017
0.36365 | -76.6885
0.36733 | -76.5712
0.371 | -76.7564
0.37467 | -76.0033
0.37835 | -76.4041
0.38202 | -76.6658
0.38569 | -76.6796
0.38937 | -76.6536
0.39304 | -77.1374
0.39671 | -77.8234
0.40039 | -77.641
0.40406 | -77.1719
0.40773 | -77.7992
0.41141 | -77.2161
0.41508 | -76.5644
0.41875 | -76.8297
0.42242 | -78.1026
0.4261 | -78.5673
0.42977 | -77.6072
0.43345 | -77.4221
0.43712 | -77.9526
0.44079 | -77.3395
0.44446 | -77.2912
0.44814 | -79.2611
0.45181 | -79.2682
0.45548 | -79.8147
0.45916 | -79.006
0.46283 | -78.37
0.4665 | -78.2555
0.47018 | -78.9023
0.47385 | -79.9811
0.47752 | -79.9634
0.4812 | -79.5586
0.48487 | -78.7649
0.48854 | -78.8276
0.49222 | -79.6283
0.49589 | -80.2551
0.49956 | -81.0329
0.50324 | -79.9316
0.50691 | -78.9127
0.51058 | -78.9889
0.51426 | -78.7638
0.51793 | -79.3002
0.5216 | -79.6553
0.52528 | -80.1131
0.52895 | -80.1911
0.53262 | -80.0363
0.5363 | -79.7739
0.53997 | -79.7524
0.54364 | -80.1287
0.54732 | -80.2976
0.55099 | -80.3283
0.55466 | -80.6106
0.55834 | -81.0276
0.56201 | -81.3077
0.56568 | -80.6628
0.56936 | -79.8909
0.57303 | -79.7379
0.5767 | -79.9039
0.58038 | -80.2478
0.58405 | -80.7025
0.58772 | -81.1429
0.5914 | -81.1095
0.59507 | -80.7178
0.59874 | -80.9505
0.60242 | -81.2641
0.60609 | -81.3822
0.60976 | -81.2393
0.61343 | -81.2084
0.61711 | -81.8997
0.62078 | -82.6379
0.62445 | -82.344
0.62813 | -81.7787
0.6318 | -81.4957
0.63548 | -81.4783
0.63915 | -81.7456
0.64282 | -82.2581
0.64649 | -82.4372
0.65017 | -82.4115
0.65384 | -82.06
0.65751 | -81.7435
0.66119 | -82.0107
0.66486 | -82.5089
0.66853 | -82.8775
0.67221 | -83.0378
0.67588 | -83.0172
0.67955 | -83.1589
0.68323 | -83.4606
0.6869 | -83.1086
0.69057 | -82.6635
0.69425 | -82.8732
0.69792 | -83.3197
0.70159 | -82.8885
0.70527 | -82.2175
0.70894 | -82.128
0.71261 | -82.4185
0.71629 | -82.6098
0.71996 | -82.4174
0.72363 | -82.4413
0.72731 | -82.7949
0.73098 | -83.0763
0.73465 | -83.2878
0.73833 | -83.7644
0.742 | -84.3319
0.74567 | -84.3974
0.74935 | -84.0013
0.75302 | -83.7455
0.75669 | -83.871
0.76036 | -83.7915
0.76404 | -83.8265
0.76771 | -84.1701
0.77138 | -84.502
0.77506 | -84.4127
0.77873 | -84.131
0.78241 | -83.7891
0.78608 | -83.9181
0.78975 | -84.6578
0.79343 | -84.9873
0.7971 | -84.8314
0.80077 | -84.7445
0.80444 | -84.764
0.80812 | -84.9866
0.81179 | -84.9621
0.81546 | -84.6404
0.81914 | -84.7037
0.82281 | -85.0616
0.82648 | -85.5052
0.83016 | -85.68
0.83383 | -85.3634
0.8375 | -85.0806
0.84118 | -85.124
0.84485 | -85.5856
0.84852 | -85.7269
0.8522 | -85.3585
0.85587 | -84.9448
0.85954 | -84.9843
0.86322 | -85.3795
0.86689 | -85.6004
0.87056 | -85.2877
0.87424 | -84.9153
0.87791 | -84.9613
0.88158 | -85.2943
0.88526 | -85.4819
0.88893 | -85.2408
0.8926 | -84.986
0.89628 | -84.8387
0.89995 | -84.7838
0.90362 | -85.3842
0.9073 | -85.7064
0.91097 | -85.6593
0.91464 | -86.2701
0.91832 | -86.9864
0.92199 | -86.8727
0.92566 | -86.0774
0.92934 | -85.5168
0.93301 | -85.398
0.93668 | -85.702
0.94036 | -86.2138
0.94403 | -85.9365
0.9477 | -85.7453
0.95137 | -86.0437
0.95505 | -86.3264
0.95872 | -86.4134
0.9624 | -86.3555
0.96607 | -86.3919
0.96974 | -86.6373
0.97342 | -87.0131
0.97709 | -87.0349
0.98076 | -86.5769
0.98443 | -86.2039
0.98811 | -86.517
0.99178 | -86.995
0.99545 | -86.9323
0.99913 | -86.6906
1.0028 | -86.8619
1.0065 | -87.3217
1.0101 | -87.3171
1.0138 | -87.2265
1.0175 | -87.2564
1.0212 | -87.252
1.0248 | -87.6016
1.0285 | -87.6751
1.0322 | -87.1414
1.0359 | -86.5404
1.0395 | -86.5948
1.0432 | -86.9984
1.0469 | -87.0236
1.0506 | -87.034
1.0542 | -86.9367
1.0579 | -86.6285
1.0616 | -86.5338
1.0652 | -86.7403
1.0689 | -87.2516
1.0726 | -87.8818
1.0763 | -88.3888
1.0799 | -88.5964
1.0836 | -88.5685
1.0873 | -88.5727
1.091 | -88.5994
1.0946 | -88.4199
1.0983 | -87.9968
1.102 | -87.5772
1.1057 | -87.4043
1.1093 | -87.5381
1.113 | -87.8681
1.1167 | -88.1899
1.1203 | -88.3401
1.124 | -88.3316
1.1277 | -88.3218
1.1314 | -88.4085
1.135 | -88.4485
1.1387 | -88.3363
1.1424 | -88.2437
1.1461 | -88.3303
1.1497 | -88.443
1.1534 | -88.3665
1.1571 | -88.2198
1.1608 | -88.2702
1.1644 | -88.5318
1.1681 | -88.7712
1.1718 | -88.8326
1.1754 | -88.7104
1.1791 | -88.4092
1.1828 | -88.0598
1.1865 | -87.8971
1.1901 | -87.9978
1.1938 | -88.2181
1.1975 | -88.3983
1.2012 | -88.4263
1.2048 | -88.3343
1.2085 | -88.3679
1.2122 | -88.5897
1.2159 | -88.6807
1.2195 | -88.5053
1.2232 | -88.4214
1.2269 | -88.6852
1.2305 | -89.1748
1.2342 | -89.557
1.2379 | -89.5987
1.2416 | -89.4085
1.2452 | -89.1445
1.2489 | -88.851
1.2526 | -88.6821
1.2563 | -88.6108
1.2599 | -88.4451
1.2636 | -88.4057
1.2673 | -88.6687
1.2709 | -88.977
1.2746 | -89.199
1.2783 | -89.4802
1.282 | -89.9214
1.2856 | -90.3824
1.2893 | -90.3214
1.293 | -89.6122
1.2967 | -88.9543
1.3003 | -88.8454
1.304 | -89.1993
1.3077 | -89.6002
1.3114 | -89.7465
1.315 | -89.6223
1.3187 | -89.3128
1.3224 | -89.0236
1.326 | -89.0617
1.3297 | -89.4758
1.3334 | -89.8476
1.3371 | -89.8755
1.3407 | -89.7642
1.3444 | -89.6694
1.3481 | -89.5979
1.3518 | -89.5202
1.3554 | -89.3718
1.3591 | -89.2139
1.3628 | -89.2825
1.3665 | -89.6453
1.3701 | -90.0416
1.3738 | -90.2409
1.3775 | -90.2407
1.3811 | -90.0433
1.3848 | -89.7232
1.3885 | -89.5479
1.3922 | -89.7298
1.3958 | -90.1389
1.3995 | -90.4498
1.4032 | -90.5797
1.4069 | -90.677
1.4105 | -90.7747
1.4142 | -90.7982
1.4179 | -90.7605
1.4216 | -90.7719
1.4252 | -90.8809
1.4289 | -90.9776
1.4326 | -91.0177
1.4362 | -91.0519
1.4399 | -90.9691
1.4436 | -90.6719
1.4473 | -90.3316
1.4509 | -90.0362
1.4546 | -89.6892
1.4583 | -89.3614
1.462 | -89.3377
1.4656 | -89.7625
1.4693 | -90.4413
1.473 | -90.9894
1.4767 | -91.2808
1.4803 | -91.3427
1.484 | -91.119
1.4877 | -90.7949
1.4913 | -90.6011
1.495 | -90.4859
1.4987 | -90.4049
1.5024 | -90.4018
1.506 | -90.3754
1.5097 | -90.2335
1.5134 | -90.0695
1.5171 | -90.0438
1.5207 | -90.3088
1.5244 | -90.827
1.5281 | -91.1331
1.5318 | -90.9114
1.5354 | -90.5625
1.5391 | -90.4036
1.5428 | -90.3781
1.5464 | -90.4587
1.5501 | -90.6978
1.5538 | -90.9393
1.5575 | -90.9431
1.5611 | -90.768
1.5648 | -90.6914
1.5685 | -90.8156
1.5722 | -91.0541
1.5758 | -91.3973
1.5795 | -91.7912
1.5832 | -91.852
1.5869 | -91.3176
1.5905 | -91.0026
1.5942 | -90.8512
1.5979 | -90.9199
1.6015 | -91.0922
1.6052 | -91.2182
1.6089 | -91.2645
1.6126 | -91.3034
1.6162 | -91.3822
1.6199 | -91.46
1.6236 | -91.4909
1.6273 | -91.4912
1.6309 | -91.4709
1.6346 | -91.4259
1.6383 | -91.3302
1.6419 | -91.2267
1.6456 | -91.1884
1.6493 | -91.2262
1.653 | -91.3041
1.6566 | -91.3554
1.6603 | -91.352
1.664 | -91.3014
1.6677 | -91.2166
1.6713 | -91.0982
1.675 | -90.9519
1.6787 | -90.8123
1.6824 | -90.712
1.686 | -90.6661
1.6897 | -90.6808
1.6934 | -90.7186
1.697 | -90.7305
1.7007 | -90.7461
1.7044 | -90.8496
1.7081 | -91.0624
1.7117 | -91.3342
1.7154 | -91.575
1.7191 | -91.6984
1.7228 | -91.6856
1.7264 | -91.6219
1.7301 | -91.6484
1.7338 | -91.8185
1.7375 | -92.1016
1.7411 | -92.4196
1.7448 | -92.658
1.7485 | -92.7788
1.7521 | -92.7889
1.7558 | -92.6792
1.7595 | -92.4595
1.7632 | -92.1963
1.7668 | -91.9969
1.7705 | -91.9038
1.7742 | -91.9264
1.7779 | -91.9814
1.7815 | -91.9389
1.7852 | -91.7571
1.7889 | -91.5701
1.7926 | -91.5444
1.7962 | -91.7242
1.7999 | -91.976
1.8036 | -92.1171
1.8072 | -92.0744
1.8109 | -91.9493
1.8146 | -91.8635
1.8183 | -91.8771
1.8219 | -91.9557
1.8256 | -92.0484
1.8293 | -92.1195
1.833 | -92.1566
1.8366 | -92.1536
1.8403 | -92.1155
1.844 | -92.0718
1.8477 | -92.0669
1.8513 | -92.1398
1.855 | -92.2571
1.8587 | -92.355
1.8623 | -92.438
1.866 | -92.5661
1.8697 | -92.7403
1.8734 | -92.8507
1.877 | -92.7725
1.8807 | -92.5467
1.8844 | -92.2953
1.8881 | -92.1111
1.8917 | -92.0401
1.8954 | -92.1253
1.8991 | -92.3984
1.9027 | -92.8419
1.9064 | -93.3728
1.9101 | -93.8304
1.9138 | -94.0189
1.9174 | -93.9097
1.9211 | -93.638
1.9248 | -93.3916
1.9285 | -93.2454
1.9321 | -93.1519
1.9358 | -93.0016
1.9395 | -92.789
1.9432 | -92.5923
1.9468 | -92.4439
1.9505 | -92.3131
1.9542 | -92.1875
1.9578 | -92.0961
1.9615 | -92.0762
1.9652 | -92.128
1.9689 | -92.2651
1.9725 | -92.4963
1.9762 | -92.8082
1.9799 | -93.1195
1.9836 | -93.3243
1.9872 | -93.3669
1.9909 | -93.2881
1.9946 | -93.197
1.9983 | -93.1822
2.0019 | -93.2537
2.0056 | -93.3869
2.0093 | -93.5545
2.0129 | -93.686
2.0166 | -93.7246
2.0203 | -93.659
2.024 | -93.4885
2.0276 | -93.2737
2.0313 | -93.1105
2.035 | -93.068
2.0387 | -93.1557
2.0423 | -93.3455
2.046 | -93.5591
2.0497 | -93.6462
2.0534 | -93.5338
2.057 | -93.2974
2.0607 | -93.0193
2.0644 | -92.7749
2.068 | -92.6192
2.0717 | -92.5899
2.0754 | -92.6703
2.0791 | -92.8408
2.0827 | -93.0939
2.0864 | -93.4078
2.0901 | -93.714
2.0938 | -93.9763
2.0974 | -94.1895
2.1011 | -94.3422
2.1048 | -94.3653
2.1085 | -94.233
2.1121 | -94.0349
2.1158 | -93.9042
2.1195 | -93.9056
2.1231 | -93.8851
2.1268 | -93.8792
2.1305 | -93.8889
2.1342 | -93.9337
2.1378 | -93.9957
2.1415 | -94.037
2.1452 | -94.0153
2.1489 | -93.8987
2.1525 | -93.7026
2.1562 | -93.5018
2.1599 | -93.3868
2.1636 | -93.3835
2.1672 | -93.4468
2.1709 | -93.4867
2.1746 | -93.4526
2.1782 | -93.3814
2.1819 | -93.3288
2.1856 | -93.3307
2.1893 | -93.3889
2.1929 | -93.5079
2.1966 | -93.6441
2.2003 | -93.7566
2.204 | -93.7889
2.2076 | -93.7236
2.2113 | -93.6104
2.215 | -93.5346
2.2187 | -93.5402
2.2223 | -93.621
2.226 | -93.7219
2.2297 | -93.7987
2.2333 | -93.818
2.237 | -93.8165
2.2407 | -93.848
2.2444 | -93.9526
2.248 | -94.097
2.2517 | -94.1777
2.2554 | -94.1191
2.2591 | -93.9417
2.2627 | -93.7445
2.2664 | -93.6643
2.2701 | -93.7644
2.2738 | -94.0341
2.2774 | -94.3476
2.2811 | -94.5514
2.2848 | -94.5999
2.2884 | -94.552
2.2921 | -94.4889
2.2958 | -94.4563
2.2995 | -94.4712
2.3031 | -94.5162
2.3068 | -94.525
2.3105 | -94.4537
2.3142 | -94.2911
2.3178 | -94.1111
2.3215 | -93.9556
2.3252 | -93.8511
2.3288 | -93.8152
2.3325 | -93.8684
2.3362 | -94.0089
2.3399 | -94.2001
2.3435 | -94.3922
2.3472 | -94.5451
2.3509 | -94.622
2.3546 | -94.6124
2.3582 | -94.5018
2.3619 | -94.3179
2.3656 | -94.1064
2.3693 | -93.9138
2.3729 | -93.7584
2.3766 | -93.6457
2.3803 | -93.5728
2.3839 | -93.5581
2.3876 | -93.6262
2.3913 | -93.7551
2.395 | -93.9033
2.3986 | -94.0259
2.4023 | -94.0947
2.406 | -94.0925
2.4097 | -94.033
2.4133 | -93.9587
2.417 | -93.9009
2.4207 | -93.8926
2.4244 | -93.9424
2.428 | -94.0533
2.4317 | -94.2159
2.4354 | -94.3849
2.439 | -94.492
2.4427 | -94.5256
2.4464 | -94.5116
2.4501 | -94.4717
2.4537 | -94.4361
2.4574 | -94.4455
2.4611 | -94.4989
2.4648 | -94.5957
2.4684 | -94.7193
2.4721 | -94.8194
2.4758 | -94.8524
2.4795 | -94.8233
2.4831 | -94.7485
2.4868 | -94.6988
2.4905 | -94.7034
2.4941 | -94.7584
2.4978 | -94.8456
2.5015 | -94.9728
2.5052 | -95.1055
2.5088 | -95.1627
2.5125 | -95.1168
2.5162 | -94.991
2.5199 | -94.8638
2.5235 | -94.7787
2.5272 | -94.7422
2.5309 | -94.7612
2.5346 | -94.8208
2.5382 | -94.9054
2.5419 | -94.9688
2.5456 | -94.9811
2.5492 | -94.9202
2.5529 | -94.8016
2.5566 | -94.6697
2.5603 | -94.5618
2.5639 | -94.4938
2.5676 | -94.4641
2.5713 | -94.4838
2.575 | -94.5542
2.5786 | -94.653
2.5823 | -94.7422
2.586 | -94.792
2.5897 | -94.8348
2.5933 | -94.9026
2.597 | -95.007
2.6007 | -95.1315
2.6043 | -95.2324
2.608 | -95.2691
2.6117 | -95.2166
2.6154 | -95.0991
2.619 | -94.9784
2.6227 | -94.8756
2.6264 | -94.7964
2.6301 | -94.7275
2.6337 | -94.707
2.6374 | -94.7798
2.6411 | -94.9352
2.6448 | -95.1367
2.6484 | -95.2945
2.6521 | -95.4396
2.6558 | -95.563
2.6594 | -95.6627
2.6631 | -95.7229
2.6668 | -95.7442
2.6705 | -95.7103
2.6741 | -95.6282
2.6778 | -95.5077
2.6815 | -95.3992
2.6852 | -95.342
2.6888 | -95.3403
2.6925 | -95.3909
2.6962 | -95.4658
2.6998 | -95.5483
2.7035 | -95.6156
2.7072 | -95.6612
2.7109 | -95.6891
2.7145 | -95.7304
2.7182 | -95.7921
2.7219 | -95.8742
2.7256 | -95.9379
2.7292 | -95.9336
2.7329 | -95.8617
2.7366 | -95.7578
2.7403 | -95.6599
2.7439 | -95.604
2.7476 | -95.5895
2.7513 | -95.6232
2.7549 | -95.6813
2.7586 | -95.7327
2.7623 | -95.7378
2.766 | -95.7169
2.7696 | -95.6871
2.7733 | -95.6808
2.777 | -95.7125
2.7807 | -95.7614
2.7843 | -95.8059
2.788 | -95.8396
2.7917 | -95.8662
2.7954 | -95.8829
2.799 | -95.8937
2.8027 | -95.902
2.8064 | -95.8912
2.81 | -95.8493
2.8137 | -95.7578
2.8174 | -95.6367
2.8211 | -95.5282
2.8247 | -95.4692
2.8284 | -95.4752
2.8321 | -95.5372
2.8358 | -95.6316
2.8394 | -95.735
2.8431 | -95.831
2.8468 | -95.9016
2.8505 | -95.9319
2.8541 | -95.9283
2.8578 | -95.8794
2.8615 | -95.7882
2.8651 | -95.6655
2.8688 | -95.5437
2.8725 | -95.4549
2.8762 | -95.4313
2.8798 | -95.4662
2.8835 | -95.5326
2.8872 | -95.6124
2.8909 | -95.6916
2.8945 | -95.7557
2.8982 | -95.7923
2.9019 | -95.8025
2.9056 | -95.8196
2.9092 | -95.8818
2.9129 | -96.01
2.9166 | -96.1837
2.9202 | -96.3747
2.9239 | -96.5478
2.9276 | -96.6713
2.9313 | -96.73
2.9349 | -96.7293
2.9386 | -96.692
2.9423 | -96.6493
2.946 | -96.624
2.9496 | -96.6041
2.9533 | -96.5831
2.957 | -96.5397
2.9606 | -96.4891
2.9643 | -96.4491
2.968 | -96.4357
2.9717 | -96.4654
2.9753 | -96.5346
2.979 | -96.6048
2.9827 | -96.6184
2.9864 | -96.5702
2.99 | -96.4625
2.9937 | -96.3385
2.9974 | -96.2449
3.0011 | -96.2128
3.0047 | -96.2528
3.0084 | -96.3392
3.0121 | -96.4326
3.0157 | -96.5066
3.0194 | -96.5588
3.0231 | -96.6085
3.0268 | -96.6751
3.0304 | -96.754
3.0341 | -96.824
3.0378 | -96.8739
3.0415 | -96.8932
3.0451 | -96.8849
3.0488 | -96.8624
3.0525 | -96.855
3.0562 | -96.8758
3.0598 | -96.9218
3.0635 | -96.9566
3.0672 | -96.9654
3.0708 | -96.9671
3.0745 | -96.966
3.0782 | -96.972
3.0819 | -96.9615
3.0855 | -96.9203
3.0892 | -96.8332
3.0929 | -96.7033
3.0966 | -96.5545
3.1002 | -96.4192
3.1039 | -96.3166
3.1076 | -96.2642
3.1113 | -96.2729
3.1149 | -96.327
3.1186 | -96.4032
3.1223 | -96.5029
3.1259 | -96.627
3.1296 | -96.7719
3.1333 | -96.9225
3.137 | -97.0519
3.1406 | -97.1296
3.1443 | -97.1441
3.148 | -97.1141
3.1517 | -97.0809
3.1553 | -97.0747
3.159 | -97.1098
3.1627 | -97.1909
3.1664 | -97.3025
3.17 | -97.3974
3.1737 | -97.4461
3.1774 | -97.3937
3.181 | -97.3534
3.1847 | -97.2815
3.1884 | -97.1792
3.1921 | -97.061
3.1957 | -96.9297
3.1994 | -96.8013
3.2031 | -96.6889
3.2068 | -96.6148
3.2104 | -96.5892
3.2141 | -96.6122
3.2178 | -96.6825
3.2215 | -96.789
3.2251 | -96.9207
3.2288 | -97.0569
3.2325 | -97.1898
3.2361 | -97.3018
3.2398 | -97.3768
3.2435 | -97.4188
3.2472 | -97.4329
3.2508 | -97.4201
3.2545 | -97.3881
3.2582 | -97.367
3.2619 | -97.3356
3.2655 | -97.287
3.2692 | -97.2253
3.2729 | -97.1474
3.2766 | -97.0718
3.2802 | -97.0051
3.2839 | -96.963
3.2876 | -96.9534
3.2912 | -96.9819
3.2949 | -97.0646
3.2986 | -97.1789
3.3023 | -97.296
3.3059 | -97.4013
3.3096 | -97.4668
3.3133 | -97.4752
3.317 | -97.455
3.3206 | -97.4182
3.3243 | -97.3741
3.328 | -97.3353
3.3316 | -97.3117
3.3353 | -97.3084
3.339 | -97.3124
3.3427 | -97.3136
3.3463 | -97.2995
3.35 | -97.2518
3.3537 | -97.1888
3.3574 | -97.1197
3.361 | -97.0576
3.3647 | -97.0244
3.3684 | -97.0234
3.3721 | -97.0596
3.3757 | -97.1214
3.3794 | -97.1968
3.3831 | -97.2818
3.3867 | -97.3725
3.3904 | -97.4718
3.3941 | -97.5716
3.3978 | -97.6518
3.4014 | -97.7082
3.4051 | -97.755
3.4088 | -97.8027
3.4125 | -97.8724
3.4161 | -97.9586
3.4198 | -98.0553
3.4235 | -98.1511
3.4272 | -98.2192
3.4308 | -98.2479
3.4345 | -98.2208
3.4382 | -98.1441
3.4418 | -98.0314
3.4455 | -97.9267
3.4492 | -97.8496
3.4529 | -97.7903
3.4565 | -97.7459
3.4602 | -97.7015
3.4639 | -97.6617
3.4676 | -97.6375
3.4712 | -97.6338
3.4749 | -97.6481
3.4786 | -97.6674
3.4823 | -97.6909
3.4859 | -97.7099
3.4896 | -97.7327
3.4933 | -97.7465
3.4969 | -97.7537
3.5006 | -97.754
3.5043 | -97.7581
3.508 | -97.7804
3.5116 | -97.809
3.5153 | -97.8279
3.519 | -97.8194
3.5227 | -97.7719
3.5263 | -97.7023
3.53 | -97.6207
3.5337 | -97.5503
3.5374 | -97.512
3.541 | -97.5284
3.5447 | -97.5975
3.5484 | -97.7007
3.552 | -97.7986
3.5557 | -97.8712
3.5594 | -97.9165
3.5631 | -97.918
3.5667 | -97.8853
3.5704 | -97.8372
3.5741 | -97.7939
3.5778 | -97.769
3.5814 | -97.7796
3.5851 | -97.8424
3.5888 | -97.9598
3.5925 | -98.0958
3.5961 | -98.2173
3.5998 | -98.3002
3.6035 | -98.3373
3.6071 | -98.3095
3.6108 | -98.2253
3.6145 | -98.1135
3.6182 | -98.0239
3.6218 | -97.9871
3.6255 | -98.0155
3.6292 | -98.1055
3.6329 | -98.2507
3.6365 | -98.4161
3.6402 | -98.5693
3.6439 | -98.6906
3.6475 | -98.7841
3.6512 | -98.8347
3.6549 | -98.8451
3.6586 | -98.831
3.6622 | -98.8056
3.6659 | -98.7805
3.6696 | -98.7738
3.6733 | -98.7617
3.6769 | -98.7393
3.6806 | -98.6941
3.6843 | -98.6231
3.688 | -98.5369
3.6916 | -98.4557
3.6953 | -98.3837
3.699 | -98.3198
3.7026 | -98.2626
3.7063 | -98.2441
3.71 | -98.2245
3.7137 | -98.2089
3.7173 | -98.2034
3.721 | -98.2282
3.7247 | -98.2525
3.7284 | -98.2797
3.732 | -98.3202
3.7357 | -98.3526
3.7394 | -98.384
3.7431 | -98.4252
3.7467 | -98.4798
3.7504 | -98.5545
3.7541 | -98.6418
3.7577 | -98.7238
3.7614 | -98.79
3.7651 | -98.8225
3.7688 | -98.8127
3.7724 | -98.756
3.7761 | -98.677
3.7798 | -98.5888
3.7835 | -98.4939
3.7871 | -98.4303
3.7908 | -98.391
3.7945 | -98.3809
3.7982 | -98.4093
3.8018 | -98.4648
3.8055 | -98.5281
3.8092 | -98.5812
3.8128 | -98.6241
3.8165 | -98.6496
3.8202 | -98.6517
3.8239 | -98.6205
3.8275 | -98.5649
3.8312 | -98.5213
3.8349 | -98.499
3.8386 | -98.4925
3.8422 | -98.5075
3.8459 | -98.548
3.8496 | -98.6048
3.8533 | -98.668
3.8569 | -98.7374
3.8606 | -98.8061
3.8643 | -98.8703
3.8679 | -98.9307
3.8716 | -98.9859
3.8753 | -99.0079
3.879 | -98.9939
3.8826 | -98.9506
3.8863 | -98.8859
3.89 | -98.8009
3.8937 | -98.6951
3.8973 | -98.5828
3.901 | -98.4704
3.9047 | -98.3655
3.9084 | -98.2798
3.912 | -98.2238
3.9157 | -98.2077
3.9194 | -98.2359
3.923 | -98.303
3.9267 | -98.4009
3.9304 | -98.514
3.9341 | -98.6313
3.9377 | -98.7421
3.9414 | -98.8364
3.9451 | -98.8953
3.9488 | -98.9215
3.9524 | -98.9209
3.9561 | -98.8905
3.9598 | -98.858
3.9635 | -98.8417
3.9671 | -98.8485
3.9708 | -98.896
3.9745 | -98.9641
3.9781 | -99.0499
3.9818 | -99.1325
3.9855 | -99.1932
3.9892 | -99.2164
3.9928 | -99.1978
3.9965 | -99.139
4.0002 | -99.0489
4.0039 | -98.9407
4.0075 | -98.8418
4.0112 | -98.7634
4.0149 | -98.7072
4.0186 | -98.6826
4.0222 | -98.679
4.0259 | -98.6852
4.0296 | -98.7015
4.0332 | -98.7112
4.0369 | -98.7172
4.0406 | -98.7105
4.0443 | -98.6896
4.0479 | -98.6646
4.0516 | -98.64
4.0553 | -98.6136
4.059 | -98.594
4.0626 | -98.5913
4.0663 | -98.6
4.07 | -98.6197
4.0736 | -98.6661
4.0773 | -98.7443
4.081 | -98.8416
4.0847 | -98.9434
4.0883 | -99.0456
4.092 | -99.1359
4.0957 | -99.2125
4.0994 | -99.2618
4.103 | -99.2914
4.1067 | -99.3071
4.1104 | -99.3104
4.1141 | -99.3086
4.1177 | -99.3244
4.1214 | -99.3632
4.1251 | -99.4254
4.1287 | -99.5155
4.1324 | -99.6028
4.1361 | -99.6678
4.1398 | -99.7026
4.1434 | -99.6851
4.1471 | -99.6188
4.1508 | -99.5145
4.1545 | -99.3897
4.1581 | -99.2813
4.1618 | -99.2074
4.1655 | -99.1612
4.1692 | -99.1489
4.1728 | -99.1633
4.1765 | -99.1747
4.1802 | -99.1698
4.1838 | -99.1392
4.1875 | -99.0829
4.1912 | -99.0063
4.1949 | -98.9317
4.1985 | -98.8675
4.2022 | -98.8306
4.2059 | -98.8392
4.2096 | -98.8976
4.2132 | -99.0059
4.2169 | -99.1519
4.2206 | -99.32
4.2243 | -99.5055
4.2279 | -99.6902
4.2316 | -99.8569
4.2353 | -99.9989
4.2389 | -100.0746
4.2426 | -100.171
4.2463 | -100.2446
4.25 | -100.2928
4.2536 | -100.3152
4.2573 | -100.3167
4.261 | -100.2879
4.2647 | -100.2437
4.2683 | -100.1864
4.272 | -100.1178
4.2757 | -100.0401
4.2794 | -99.9535
4.283 | -99.8934
4.2867 | -99.8585
4.2904 | -99.8418
4.294 | -99.8501
4.2977 | -99.8739
4.3014 | -99.9022
4.3051 | -99.9341
4.3087 | -99.9559
4.3124 | -99.9577
4.3161 | -99.9442
4.3198 | -99.9178
4.3234 | -99.8896
4.3271 | -99.8685
4.3308 | -99.8534
4.3345 | -99.8364
4.3381 | -99.8129
4.3418 | -99.7904
4.3455 | -99.7587
4.3491 | -99.7212
4.3528 | -99.6662
4.3565 | -99.6121
4.3602 | -99.564
4.3638 | -99.5217
4.3675 | -99.4911
4.3712 | -99.4745
4.3749 | -99.4652
4.3785 | -99.4638
4.3822 | -99.4693
4.3859 | -99.482
4.3895 | -99.5025
4.3932 | -99.5134
4.3969 | -99.5032
4.4006 | -99.4716
4.4042 | -99.4167
4.4079 | -99.3398
4.4116 | -99.2569
4.4153 | -99.1801
4.4189 | -99.1205
4.4226 | -99.0924
4.4263 | -99.0989
4.43 | -99.1446
4.4336 | -99.2125
4.4373 | -99.2953
4.441 | -99.3848
4.4446 | -99.4769
4.4483 | -99.5596
4.452 | -99.6269
4.4557 | -99.6722
4.4593 | -99.703
4.463 | -99.7331
4.4667 | -99.7542
4.4704 | -99.7654
4.474 | -99.775
4.4777 | -99.7822
4.4814 | -99.7724
4.4851 | -99.7628
4.4887 | -99.7469
4.4924 | -99.7269
4.4961 | -99.7219
4.4997 | -99.7321
4.5034 | -99.7723
4.5071 | -99.8341
4.5108 | -99.9084
4.5144 | -99.9916
4.5181 | -100.0854
4.5218 | -100.1675
4.5255 | -100.2175
4.5291 | -100.2279
4.5328 | -100.1966
4.5365 | -100.1359
4.5402 | -100.0579
4.5438 | -99.9684
4.5475 | -99.889
4.5512 | -99.8307
4.5548 | -99.7911
4.5585 | -99.7758
4.5622 | -99.7834
4.5659 | -99.8059
4.5695 | -99.8472
4.5732 | -99.8972
4.5769 | -99.9469
4.5806 | -99.997
4.5842 | -100.0321
4.5879 | -100.052
4.5916 | -100.0647
4.5953 | -100.0703
4.5989 | -100.0634
4.6026 | -100.0462
4.6063 | -100.0247
4.6099 | -99.9888
4.6136 | -99.9332
4.6173 | -99.8682
4.621 | -99.8009
4.6246 | -99.7442
4.6283 | -99.7046
4.632 | -99.6883
4.6357 | -99.7056
4.6393 | -99.7621
4.643 | -99.8453
4.6467 | -99.9613
4.6504 | -100.0962
4.654 | -100.2424
4.6577 | -100.3873
4.6614 | -100.5244
4.665 | -100.6625
4.6687 | -100.7944
4.6724 | -100.9028
4.6761 | -100.986
4.6797 | -101.0359
4.6834 | -101.0592
4.6871 | -101.048
4.6908 | -100.9928
4.6944 | -100.9034
4.6981 | -100.7956
4.7018 | -100.6849
4.7054 | -100.5734
4.7091 | -100.4781
4.7128 | -100.414
4.7165 | -100.3644
4.7201 | -100.3368
4.7238 | -100.3237
4.7275 | -100.3185
4.7312 | -100.315
4.7348 | -100.3044
4.7385 | -100.2917
4.7422 | -100.2829
4.7459 | -100.2743
4.7495 | -100.2796
4.7532 | -100.283
4.7569 | -100.2863
4.7605 | -100.291
4.7642 | -100.2992
4.7679 | -100.3227
4.7716 | -100.3476
4.7752 | -100.383
4.7789 | -100.4379
4.7826 | -100.5119
4.7863 | -100.6072
4.7899 | -100.7215
4.7936 | -100.8447
4.7973 | -100.969
4.801 | -101.0802
4.8046 | -101.1802
4.8083 | -101.2538
4.812 | -101.2929
4.8156 | -101.301
4.8193 | -101.2676
4.823 | -101.1935
4.8267 | -101.0898
4.8303 | -100.9682
4.834 | -100.8439
4.8377 | -100.7162
4.8414 | -100.6043
4.845 | -100.5127
4.8487 | -100.4484
4.8524 | -100.4084
4.8561 | -100.3946
4.8597 | -100.3951
4.8634 | -100.4106
4.8671 | -100.4403
4.8707 | -100.4779
4.8744 | -100.5156
4.8781 | -100.5453
4.8818 | -100.5578
4.8854 | -100.5616
4.8891 | -100.558
4.8928 | -100.5537
4.8965 | -100.5607
4.9001 | -100.5759
4.9038 | -100.6111
4.9075 | -100.6577
4.9112 | -100.7006
4.9148 | -100.7407
4.9185 | -100.769
4.9222 | -100.7765
4.9258 | -100.7547
4.9295 | -100.7131
4.9332 | -100.6572
4.9369 | -100.6028
4.9405 | -100.5554
4.9442 | -100.5283
4.9479 | -100.524
4.9516 | -100.5459
4.9552 | -100.5892
4.9589 | -100.6525
4.9626 | -100.7211
4.9663 | -100.7849
4.9699 | -100.8357
4.9736 | -100.8681
4.9773 | -100.8799
4.9809 | -100.862
4.9846 | -100.8247
4.9883 | -100.7812
4.992 | -100.7266
4.9956 | -100.6688
4.9993 | -100.6221
5.003 | -100.598
5.0067 | -100.5918
5.0103 | -100.6011
5.014 | -100.6205
5.0177 | -100.6482
5.0214 | -100.6794
5.025 | -100.7096
5.0287 | -100.7424
5.0324 | -100.7757
5.036 | -100.801
5.0397 | -100.824
5.0434 | -100.8334
5.0471 | -100.8307
5.0507 | -100.822
5.0544 | -100.7996
5.0581 | -100.7815
5.0618 | -100.7645
5.0654 | -100.7531
5.0691 | -100.7618
5.0728 | -100.7843
5.0765 | -100.8172
5.0801 | -100.8512
5.0838 | -100.8795
5.0875 | -100.9042
5.0911 | -100.916
5.0948 | -100.9095
5.0985 | -100.8771
5.1022 | -100.8151
5.1058 | -100.7312
5.1095 | -100.6345
5.1132 | -100.5438
5.1169 | -100.4673
5.1205 | -100.4193
5.1242 | -100.4021
5.1279 | -100.418
5.1315 | -100.4684
5.1352 | -100.5411
5.1389 | -100.6338
5.1426 | -100.7389
5.1462 | -100.8408
5.1499 | -100.94
5.1536 | -101.0403
5.1573 | -101.132
5.1609 | -101.2165
5.1646 | -101.2885
5.1683 | -101.3501
5.172 | -101.3981
5.1756 | -101.4364
5.1793 | -101.4706
5.183 | -101.4857
5.1866 | -101.482
5.1903 | -101.4587
5.194 | -101.4242
5.1977 | -101.3779
5.2013 | -101.3251
5.205 | -101.2666
5.2087 | -101.2052
5.2124 | -101.1418
5.216 | -101.0813
5.2197 | -101.0231
5.2234 | -100.9642
5.2271 | -100.9077
5.2307 | -100.8534
5.2344 | -100.8051
5.2381 | -100.7631
5.2417 | -100.7336
5.2454 | -100.721
5.2491 | -100.7221
5.2528 | -100.7418
5.2564 | -100.7805
5.2601 | -100.8363
5.2638 | -100.9067
5.2675 | -100.9858
5.2711 | -101.0747
5.2748 | -101.1741
5.2785 | -101.2652
5.2822 | -101.3462
5.2858 | -101.4192
5.2895 | -101.4686
5.2932 | -101.4928
5.2968 | -101.4453
5.3005 | -101.4184
5.3042 | -101.3794
5.3079 | -101.3285
5.3115 | -101.2651
5.3152 | -101.2032
5.3189 | -101.1519
5.3226 | -101.1113
5.3262 | -101.0792
5.3299 | -101.0603
5.3336 | -101.0563
5.3373 | -101.0634
5.3409 | -101.0857
5.3446 | -101.1071
5.3483 | -101.1218
5.3519 | -101.1252
5.3556 | -101.1218
5.3593 | -101.1171
5.363 | -101.1123
5.3666 | -101.108
5.3703 | -101.1125
5.374 | -101.1237
5.3777 | -101.1447
5.3813 | -101.1655
5.385 | -101.1969
5.3887 | -101.2307
5.3924 | -101.2628
5.396 | -101.2964
5.3997 | -101.3312
5.4034 | -101.3599
5.407 | -101.3807
5.4107 | -101.4031
5.4144 | -101.4222
5.4181 | -101.4304
5.4217 | -101.4287
5.4254 | -101.418
5.4291 | -101.4132
5.4328 | -101.4134
5.4364 | -101.411
5.4401 | -101.4079
5.4438 | -101.4125
5.4474 | -101.4201
5.4511 | -101.4255
5.4548 | -101.4386
5.4585 | -101.4585
5.4621 | -101.4825
5.4658 | -101.5068
5.4695 | -101.5332
5.4732 | -101.5587
5.4768 | -101.5887
5.4805 | -101.6242
5.4842 | -101.6634
5.4879 | -101.7042
5.4915 | -101.7462
5.4952 | -101.7822
5.4989 | -101.8113
5.5025 | -101.8268
5.5062 | -101.8346
5.5099 | -101.8278
5.5136 | -101.8036
5.5172 | -101.7594
5.5209 | -101.7042
5.5246 | -101.6283
5.5283 | -101.5406
5.5319 | -101.4469
5.5356 | -101.352
5.5393 | -101.2589
5.543 | -101.1788
5.5466 | -101.1258
5.5503 | -101.102
5.554 | -101.0943
5.5576 | -101.1022
5.5613 | -101.1283
5.565 | -101.1704
5.5687 | -101.2177
5.5723 | -101.2706
5.576 | -101.3301
5.5797 | -101.3871
5.5834 | -101.4369
5.587 | -101.4859
5.5907 | -101.5312
5.5944 | -101.5753
5.5981 | -101.6128
5.6017 | -101.6543
5.6054 | -101.6888
5.6091 | -101.7222
5.6127 | -101.7457
5.6164 | -101.7619
5.6201 | -101.7693
5.6238 | -101.7584
5.6274 | -101.7326
5.6311 | -101.6907
5.6348 | -101.6465
5.6385 | -101.608
5.6421 | -101.5824
5.6458 | -101.564
5.6495 | -101.557
5.6532 | -101.5756
5.6568 | -101.6153
5.6605 | -101.6716
5.6642 | -101.7498
5.6678 | -101.8364
5.6715 | -101.9157
5.6752 | -101.986
5.6789 | -102.0362
5.6825 | -102.0674
5.6862 | -102.0702
5.6899 | -102.0422
5.6936 | -101.9892
5.6972 | -101.9158
5.7009 | -101.8305
5.7046 | -101.7438
5.7083 | -101.6655
5.7119 | -101.6008
5.7156 | -101.5583
5.7193 | -101.541
5.7229 | -101.5477
5.7266 | -101.5727
5.7303 | -101.6112
5.734 | -101.6591
5.7376 | -101.7091
5.7413 | -101.7556
5.745 | -101.7979
5.7487 | -101.8331
5.7523 | -101.8677
5.756 | -101.9072
5.7597 | -101.9402
5.7633 | -101.9673
5.767 | -101.9841
5.7707 | -101.9883
5.7744 | -101.9872
5.778 | -101.9714
5.7817 | -101.9384
5.7854 | -101.8853
5.7891 | -101.8328
5.7927 | -101.7797
5.7964 | -101.7138
5.8001 | -101.6395
5.8038 | -101.5637
5.8074 | -101.4867
5.8111 | -101.4118
5.8148 | -101.3476
5.8184 | -101.2951
5.8221 | -101.256
5.8258 | -101.266
5.8295 | -101.2631
5.8331 | -101.2778
5.8368 | -101.3013
5.8405 | -101.3347
5.8442 | -101.3715
5.8478 | -101.4118
5.8515 | -101.4537
5.8552 | -101.5
5.8589 | -101.5517
5.8625 | -101.6064
5.8662 | -101.6666
5.8699 | -101.7348
5.8735 | -101.8103
5.8772 | -101.8751
5.8809 | -101.9424
5.8846 | -102.0158
5.8882 | -102.0767
5.8919 | -102.1265
5.8956 | -102.1646
5.8993 | -102.1935
5.9029 | -102.2071
5.9066 | -102.206
5.9103 | -102.1973
5.914 | -102.1812
5.9176 | -102.1637
5.9213 | -102.1361
5.925 | -102.1027
5.9286 | -102.0706
5.9323 | -102.0354
5.936 | -102.0047
5.9397 | -101.9856
5.9433 | -101.9693
5.947 | -101.9554
5.9507 | -101.9487
5.9544 | -101.949
5.958 | -101.9467
5.9617 | -101.9442
5.9654 | -101.9405
5.9691 | -101.9356
5.9727 | -101.9233
5.9764 | -101.9066
5.9801 | -101.8866
5.9837 | -101.8563
5.9874 | -101.8246
5.9911 | -101.7904
5.9948 | -101.7573
5.9984 | -101.733
6.0021 | -101.7136
6.0058 | -101.6982
6.0095 | -101.6901
6.0131 | -101.6916
6.0168 | -101.7009
6.0205 | -101.7161
6.0242 | -101.734
6.0278 | -101.7499
6.0315 | -101.7701
6.0352 | -101.8004
6.0388 | -101.8295
6.0425 | -101.8621
6.0462 | -101.895
6.0499 | -101.9221
6.0535 | -101.9376
6.0572 | -101.9495
6.0609 | -101.9558
6.0646 | -101.9486
6.0682 | -101.9319
6.0719 | -101.9164
6.0756 | -101.9024
6.0793 | -101.886
6.0829 | -101.8711
6.0866 | -101.8648
6.0903 | -101.8589
6.0939 | -101.8632
6.0976 | -101.8709
6.1013 | -101.8873
6.105 | -101.9035
6.1086 | -101.9193
6.1123 | -101.937
6.116 | -101.9505
6.1197 | -101.9568
6.1233 | -101.9494
6.127 | -101.9318
6.1307 | -101.9113
6.1344 | -101.8839
6.138 | -101.8551
6.1417 | -101.8311
6.1454 | -101.8168
6.149 | -101.7968
6.1527 | -101.7848
6.1564 | -101.7831
6.1601 | -101.7869
6.1637 | -101.7933
6.1674 | -101.8017
6.1711 | -101.8112
6.1748 | -101.8225
6.1784 | -101.8336
6.1821 | -101.8479
6.1858 | -101.8577
6.1894 | -101.8669
6.1931 | -101.8725
6.1968 | -101.8788
6.2005 | -101.8803
6.2041 | -101.8762
6.2078 | -101.871
6.2115 | -101.8645
6.2152 | -101.8563
6.2188 | -101.8494
6.2225 | -101.8481
6.2262 | -101.8529
6.2299 | -101.8662
6.2335 | -101.8895
6.2372 | -101.917
6.2409 | -101.9604
6.2445 | -102.0113
6.2482 | -102.0685
6.2519 | -102.1218
6.2556 | -102.1628
6.2592 | -102.1998
6.2629 | -102.2283
6.2666 | -102.2488
6.2703 | -102.2662
6.2739 | -102.2785
6.2776 | -102.2879
6.2813 | -102.2939
6.285 | -102.3001
6.2886 | -102.3057
6.2923 | -102.2979
6.296 | -102.2815
6.2996 | -102.2593
6.3033 | -102.2257
6.307 | -102.1866
6.3107 | -102.1375
6.3143 | -102.0734
6.318 | -101.9998
6.3217 | -101.9219
6.3254 | -101.8474
6.329 | -101.775
6.3327 | -101.714
6.3364 | -101.6774
6.3401 | -101.6627
6.3437 | -101.6754
6.3474 | -101.7124
6.3511 | -101.7666
6.3547 | -101.8395
6.3584 | -101.9332
6.3621 | -102.0076
6.3658 | -102.0784
6.3694 | -102.1507
6.3731 | -102.2168
6.3768 | -102.2737
6.3805 | -102.3172
6.3841 | -102.3524
6.3878 | -102.388
6.3915 | -102.4141
6.3952 | -102.428
6.3988 | -102.4359
6.4025 | -102.442
6.4062 | -102.4522
6.4098 | -102.4613
6.4135 | -102.4777
6.4172 | -102.4967
6.4209 | -102.5192
6.4245 | -102.5488
6.4282 | -102.5749
6.4319 | -102.593
6.4356 | -102.6115
6.4392 | -102.6257
6.4429 | -102.6414
6.4466 | -102.6559
6.4502 | -102.6668
6.4539 | -102.6798
6.4576 | -102.6926
6.4613 | -102.7023
6.4649 | -102.72
6.4686 | -102.738
6.4723 | -102.7585
6.476 | -102.7773
6.4796 | -102.7935
6.4833 | -102.814
6.487 | -102.8343
6.4907 | -102.8526
6.4943 | -102.8677
6.498 | -102.8854
6.5017 | -102.8978
6.5053 | -102.9044
6.509 | -102.9134
6.5127 | -102.9145
6.5164 | -102.9159
6.52 | -102.9141
6.5237 | -102.9081
6.5274 | -102.902
6.5311 | -102.8951
6.5347 | -102.8862
6.5384 | -102.877
6.5421 | -102.8595
6.5458 | -102.838
6.5494 | -102.8169
6.5531 | -102.7853
6.5568 | -102.7531
6.5604 | -102.7206
6.5641 | -102.6822
6.5678 | -102.6476
6.5715 | -102.6198
6.5751 | -102.6004
6.5788 | -102.5946
6.5825 | -102.6018
6.5862 | -102.625
6.5898 | -102.66
6.5935 | -102.7121
6.5972 | -102.7746
6.6009 | -102.8447
6.6045 | -102.9249
6.6082 | -103.0052
6.6119 | -103.0835
6.6155 | -103.1543
6.6192 | -103.2127
6.6229 | -103.2578
6.6266 | -103.2883
6.6302 | -103.3004
6.6339 | -103.2976
6.6376 | -103.2711
6.6413 | -103.2309
6.6449 | -103.1815
6.6486 | -103.1297
6.6523 | -103.0753
6.656 | -103.0181
6.6596 | -102.9603
6.6633 | -102.9051
6.667 | -102.8508
6.6706 | -102.807
6.6743 | -102.7734
6.678 | -102.7479
6.6817 | -102.7261
6.6853 | -102.7143
6.689 | -102.7104
6.6927 | -102.7097
6.6964 | -102.7084
6.7 | -102.7089
6.7037 | -102.7031
6.7074 | -102.7007
6.7111 | -102.7003
6.7147 | -102.6989
6.7184 | -102.697
6.7221 | -102.7019
6.7257 | -102.7096
6.7294 | -102.7177
6.7331 | -102.7312
6.7368 | -102.7557
6.7404 | -102.7845
6.7441 | -102.8086
6.7478 | -102.8301
6.7515 | -102.8528
6.7551 | -102.8707
6.7588 | -102.8861
6.7625 | -102.8958
6.7662 | -102.8972
6.7698 | -102.8876
6.7735 | -102.877
6.7772 | -102.8644
6.7808 | -102.8537
6.7845 | -102.8349
6.7882 | -102.82
6.7919 | -102.8108
6.7955 | -102.8068
6.7992 | -102.7986
6.8029 | -102.7997
6.8066 | -102.7999
6.8102 | -102.7945
6.8139 | -102.7907
6.8176 | -102.7921
6.8213 | -102.7951
6.8249 | -102.7935
6.8286 | -102.7967
6.8323 | -102.8049
6.8359 | -102.817
6.8396 | -102.8317
6.8433 | -102.8503
6.847 | -102.8736
6.8506 | -102.8983
6.8543 | -102.921
6.858 | -102.9302
6.8617 | -102.9281
6.8653 | -102.9224
6.869 | -102.9063
6.8727 | -102.8757
6.8763 | -102.8372
6.88 | -102.7977
6.8837 | -102.7655
6.8874 | -102.7531
6.891 | -102.7375
6.8947 | -102.7196
6.8984 | -102.7113
6.9021 | -102.7169
6.9057 | -102.7252
6.9094 | -102.731
6.9131 | -102.7386
6.9168 | -102.7435
6.9204 | -102.7414
6.9241 | -102.7269
6.9278 | -102.7137
6.9314 | -102.6949
6.9351 | -102.6651
6.9388 | -102.636
6.9425 | -102.6065
6.9461 | -102.577
6.9498 | -102.5506
6.9535 | -102.5323
6.9572 | -102.5256
6.9608 | -102.5338
6.9645 | -102.5553
6.9682 | -102.5843
6.9719 | -102.6204
6.9755 | -102.6649
6.9792 | -102.7186
6.9829 | -102.777
6.9865 | -102.8345
6.9902 | -102.8897
6.9939 | -102.9353
6.9976 | -102.9774
7.0012 | -103.0156
7.0049 | -103.0393
7.0086 | -103.0504
7.0123 | -103.0489
7.0159 | -103.0535
7.0196 | -103.0485
7.0233 | -103.0352
7.027 | -103.0166
7.0306 | -103.0034
7.0343 | -102.9871
7.038 | -102.9673
7.0416 | -102.9507
7.0453 | -102.9307
7.049 | -102.9172
7.0527 | -102.9107
7.0563 | -102.9083
7.06 | -102.9132
7.0637 | -102.9164
7.0674 | -102.9274
7.071 | -102.9453
7.0747 | -102.9687
7.0784 | -102.9993
7.0821 | -103.0307
7.0857 | -103.0633
7.0894 | -103.102
7.0931 | -103.1452
7.0967 | -103.1894
7.1004 | -103.2367
7.1041 | -103.2861
7.1078 | -103.3345
7.1114 | -103.3846
7.1151 | -103.4298
7.1188 | -103.4711
7.1225 | -103.5017
7.1261 | -103.5271
7.1298 | -103.5374
7.1335 | -103.5341
7.1372 | -103.5169
7.1408 | -103.4879
7.1445 | -103.4553
7.1482 | -103.4147
7.1518 | -103.3675
7.1555 | -103.312
7.1592 | -103.2529
7.1629 | -103.1952
7.1665 | -103.1465
7.1702 | -103.105
7.1739 | -103.0724
7.1776 | -103.0517
7.1812 | -103.0464
7.1849 | -103.0531
7.1886 | -103.0681
7.1922 | -103.0863
7.1959 | -103.104
7.1996 | -103.1239
7.2033 | -103.1448
7.2069 | -103.1624
7.2106 | -103.1746
7.2143 | -103.1802
7.218 | -103.1832
7.2216 | -103.1882
7.2253 | -103.1913
7.229 | -103.1872
7.2327 | -103.1845
7.2363 | -103.1879
7.24 | -103.1889
7.2437 | -103.1796
7.2473 | -103.1645
7.251 | -103.1474
7.2547 | -103.1335
7.2584 | -103.1233
7.262 | -103.1182
7.2657 | -103.1188
7.2694 | -103.1222
7.2731 | -103.1293
7.2767 | -103.1424
7.2804 | -103.1545
7.2841 | -103.1745
7.2878 | -103.1979
7.2914 | -103.2271
7.2951 | -103.2647
7.2988 | -103.3062
7.3024 | -103.3478
7.3061 | -103.3892
7.3098 | -103.4308
7.3135 | -103.4689
7.3171 | -103.5033
7.3208 | -103.5271
7.3245 | -103.5336
7.3282 | -103.5249
7.3318 | -103.5041
7.3355 | -103.4761
7.3392 | -103.4386
7.3429 | -103.3925
7.3465 | -103.3452
7.3502 | -103.297
7.3539 | -103.2557
7.3575 | -103.2095
7.3612 | -103.1605
7.3649 | -103.1156
7.3686 | -103.0741
7.3722 | -103.0386
7.3759 | -103.0137
7.3796 | -102.9951
7.3833 | -102.9868
7.3869 | -102.9824
7.3906 | -102.9749
7.3943 | -102.9717
7.398 | -102.9767
7.4016 | -102.9855
7.4053 | -102.9973
7.409 | -103.0134
7.4126 | -103.0298
7.4163 | -103.0532
7.42 | -103.0692
7.4237 | -103.0823
7.4273 | -103.0976
7.431 | -103.1193
7.4347 | -103.143
7.4384 | -103.1662
7.442 | -103.1814
7.4457 | -103.1901
7.4494 | -103.1975
7.4531 | -103.1993
7.4567 | -103.1962
7.4604 | -103.1945
7.4641 | -103.1862
7.4677 | -103.1791
7.4714 | -103.1751
7.4751 | -103.1695
7.4788 | -103.1617
7.4824 | -103.1546
7.4861 | -103.1521
7.4898 | -103.1598
7.4935 | -103.1686
7.4971 | -103.1781
7.5008 | -103.1948
7.5045 | -103.212
7.5081 | -103.2253
7.5118 | -103.2337
7.5155 | -103.2394
7.5192 | -103.2427
7.5228 | -103.2481
7.5265 | -103.2477
7.5302 | -103.242
7.5339 | -103.2372
7.5375 | -103.2307
7.5412 | -103.2249
7.5449 | -103.2211
7.5486 | -103.2206
7.5522 | -103.2225
7.5559 | -103.2328
7.5596 | -103.2422
7.5632 | -103.253
7.5669 | -103.2702
7.5706 | -103.2878
7.5743 | -103.3034
7.5779 | -103.3149
7.5816 | -103.3297
7.5853 | -103.348
7.589 | -103.3705
7.5926 | -103.3999
7.5963 | -103.4303
7.6 | -103.4573
7.6037 | -103.4877
7.6073 | -103.5241
7.611 | -103.5619
7.6147 | -103.5933
7.6183 | -103.6209
7.622 | -103.6428
7.6257 | -103.6674
7.6294 | -103.6826
7.633 | -103.6901
7.6367 | -103.6857
7.6404 | -103.6816
7.6441 | -103.6703
7.6477 | -103.652
7.6514 | -103.6354
7.6551 | -103.6133
7.6588 | -103.5805
7.6624 | -103.5466
7.6661 | -103.5163
7.6698 | -103.487
7.6734 | -103.4634
7.6771 | -103.4345
7.6808 | -103.4025
7.6845 | -103.3755
7.6881 | -103.3516
7.6918 | -103.3388
7.6955 | -103.3231
7.6992 | -103.3091
7.7028 | -103.2962
7.7065 | -103.2843
7.7102 | -103.275
7.7139 | -103.2702
7.7175 | -103.2667
7.7212 | -103.2685
7.7249 | -103.2754
7.7285 | -103.2859
7.7322 | -103.2986
7.7359 | -103.3181
7.7396 | -103.3388
7.7432 | -103.3682
7.7469 | -103.4044
7.7506 | -103.4481
7.7543 | -103.4944
7.7579 | -103.5492
7.7616 | -103.5977
7.7653 | -103.6477
7.769 | -103.7011
7.7726 | -103.7581
7.7763 | -103.81
7.78 | -103.8548
7.7836 | -103.8914
7.7873 | -103.9183
7.791 | -103.9387
7.7947 | -103.9477
7.7983 | -103.9465
7.802 | -103.9367
7.8057 | -103.9097
7.8094 | -103.8757
7.813 | -103.8347
7.8167 | -103.7951
7.8204 | -103.7506
7.8241 | -103.7049
7.8277 | -103.6626
7.8314 | -103.6231
7.8351 | -103.5867
7.8387 | -103.5546
7.8424 | -103.5322
7.8461 | -103.5188
7.8498 | -103.5153
7.8534 | -103.5165
7.8571 | -103.524
7.8608 | -103.5356
7.8645 | -103.5523
7.8681 | -103.5731
7.8718 | -103.5978
7.8755 | -103.6241
7.8792 | -103.6458
7.8828 | -103.6598
7.8865 | -103.6731
7.8902 | -103.6874
7.8938 | -103.697
7.8975 | -103.7076
7.9012 | -103.7175
7.9049 | -103.726
7.9085 | -103.7393
7.9122 | -103.758
7.9159 | -103.7776
7.9196 | -103.8024
7.9232 | -103.8315
7.9269 | -103.8563
7.9306 | -103.8786
7.9342 | -103.9049
7.9379 | -103.9268
7.9416 | -103.9445
7.9453 | -103.9572
7.9489 | -103.9548
7.9526 | -103.9613
7.9563 | -103.9692
7.96 | -103.973
7.9636 | -103.9737
7.9673 | -103.9698
7.971 | -103.9663
7.9747 | -103.9611
7.9783 | -103.9578
7.982 | -103.9509
7.9857 | -103.9465
7.9893 | -103.9409
7.993 | -103.9433
7.9967 | -103.9468
8.0004 | -103.9534
8.004 | -103.9606
8.0077 | -103.9657
8.0114 | -103.9767
8.0151 | -103.9898
8.0187 | -103.9983
8.0224 | -104.0097
8.0261 | -104.0193
8.0298 | -104.0184
8.0334 | -104.0111
8.0371 | -103.9967
8.0408 | -103.9784
8.0444 | -103.9599
8.0481 | -103.9321
8.0518 | -103.9004
8.0555 | -103.8633
8.0591 | -103.8228
8.0628 | -103.7819
8.0665 | -103.7435
8.0702 | -103.7079
8.0738 | -103.6753
8.0775 | -103.6488
8.0812 | -103.6262
8.0849 | -103.6103
8.0885 | -103.601
8.0922 | -103.6034
8.0959 | -103.6095
8.0995 | -103.6172
8.1032 | -103.6336
8.1069 | -103.6519
8.1106 | -103.6772
8.1142 | -103.7032
8.1179 | -103.7306
8.1216 | -103.7595
8.1253 | -103.7883
8.1289 | -103.814
8.1326 | -103.8339
8.1363 | -103.8553
8.14 | -103.879
8.1436 | -103.8981
8.1473 | -103.9114
8.151 | -103.9216
8.1546 | -103.9374
8.1583 | -103.9503
8.162 | -103.9579
8.1657 | -103.9612
8.1693 | -103.9641
8.173 | -103.9633
8.1767 | -103.9697
8.1804 | -103.9743
8.184 | -103.9785
8.1877 | -103.9798
8.1914 | -103.9869
8.1951 | -103.9989
8.1987 | -104.0143
8.2024 | -104.0358
8.2061 | -104.056
8.2097 | -104.0757
8.2134 | -104.0979
8.2171 | -104.1223
8.2208 | -104.142
8.2244 | -104.1586
8.2281 | -104.1725
8.2318 | -104.1786
8.2355 | -104.1857
8.2391 | -104.1888
8.2428 | -104.188
8.2465 | -104.1875
8.2501 | -104.1842
8.2538 | -104.1789
8.2575 | -104.1721
8.2612 | -104.168
8.2648 | -104.162
8.2685 | -104.1563
8.2722 | -104.1515
8.2759 | -104.14
8.2795 | -104.128
8.2832 | -104.1182
8.2869 | -104.1044
8.2906 | -104.091
8.2942 | -104.0762
8.2979 | -104.0533
8.3016 | -104.0325
8.3052 | -104.0106
8.3089 | -103.9875
8.3126 | -103.9571
8.3163 | -103.9254
8.3199 | -103.8973
8.3236 | -103.8742
8.3273 | -103.8621
8.331 | -103.8562
8.3346 | -103.8511
8.3383 | -103.8528
8.342 | -103.867
8.3457 | -103.8891
8.3493 | -103.9121
8.353 | -103.9394
8.3567 | -103.9692
8.3603 | -104.0004
8.364 | -104.0331
8.3677 | -104.0599
8.3714 | -104.0826
8.375 | -104.0993
8.3787 | -104.1064
8.3824 | -104.106
8.3861 | -104.1038
8.3897 | -104.0962
8.3934 | -104.0851
8.3971 | -104.074
8.4008 | -104.0567
8.4044 | -104.031
8.4081 | -104.0095
8.4118 | -103.9867
8.4154 | -103.959
8.4191 | -103.9324
8.4228 | -103.9046
8.4265 | -103.8786
8.4301 | -103.8503
8.4338 | -103.8247
8.4375 | -103.8053
8.4412 | -103.7828
8.4448 | -103.7548
8.4485 | -103.7299
8.4522 | -103.7042
8.4559 | -103.6836
8.4595 | -103.6709
8.4632 | -103.6608
8.4669 | -103.654
8.4705 | -103.6566
8.4742 | -103.6664
8.4779 | -103.7034
8.4816 | -103.7281
8.4852 | -103.7594
8.4889 | -103.7931
8.4926 | -103.8378
8.4963 | -103.8925
8.4999 | -103.9447
8.5036 | -103.9975
8.5073 | -104.0574
8.511 | -104.1166
8.5146 | -104.173
8.5183 | -104.2236
8.522 | -104.2666
8.5256 | -104.3051
8.5293 | -104.3392
8.533 | -104.3643
8.5367 | -104.3837
8.5403 | -104.4021
8.544 | -104.4156
8.5477 | -104.4255
8.5514 | -104.4309
8.555 | -104.4359
8.5587 | -104.4347
8.5624 | -104.4273
8.5661 | -104.4175
8.5697 | -104.4125
8.5734 | -104.4044
8.5771 | -104.3966
8.5807 | -104.3833
8.5844 | -104.3674
8.5881 | -104.3547
8.5918 | -104.3381
8.5954 | -104.3187
8.5991 | -104.3019
8.6028 | -104.2851
8.6065 | -104.265
8.6101 | -104.245
8.6138 | -104.2273
8.6175 | -104.2106
8.6211 | -104.1928
8.6248 | -104.1779
8.6285 | -104.1717
8.6322 | -104.1655
8.6358 | -104.1576
8.6395 | -104.1544
8.6432 | -104.1552
8.6469 | -104.1598
8.6505 | -104.1707
8.6542 | -104.187
8.6579 | -104.202
8.6616 | -104.2166
8.6652 | -104.2358
8.6689 | -104.2573
8.6726 | -104.277
8.6762 | -104.2927
8.6799 | -104.3107
8.6836 | -104.3287
8.6873 | -104.3475
8.6909 | -104.3623
8.6946 | -104.3792
8.6983 | -104.3898
8.702 | -104.3934
8.7056 | -104.3975
8.7093 | -104.4006
8.713 | -104.3971
8.7167 | -104.3894
8.7203 | -104.3831
8.724 | -104.3726
8.7277 | -104.3545
8.7313 | -104.3393
8.735 | -104.3215
8.7387 | -104.2993
8.7424 | -104.2768
8.746 | -104.2577
8.7497 | -104.2397
8.7534 | -104.2196
8.7571 | -104.2008
8.7607 | -104.1912
8.7644 | -104.1816
8.7681 | -104.1755
8.7718 | -104.1763
8.7754 | -104.1801
8.7791 | -104.1825
8.7828 | -104.1868
8.7864 | -104.2001
8.7901 | -104.2163
8.7938 | -104.2267
8.7975 | -104.2372
8.8011 | -104.2523
8.8048 | -104.272
8.8085 | -104.2888
8.8122 | -104.3043
8.8158 | -104.3182
8.8195 | -104.3311
8.8232 | -104.341
8.8269 | -104.3524
8.8305 | -104.3651
8.8342 | -104.3748
8.8379 | -104.3853
8.8415 | -104.3958
8.8452 | -104.4039
8.8489 | -104.4112
8.8526 | -104.4231
8.8562 | -104.4289
8.8599 | -104.4317
8.8636 | -104.4379
8.8673 | -104.4457
8.8709 | -104.454
8.8746 | -104.4633
8.8783 | -104.4711
8.882 | -104.4821
8.8856 | -104.4967
8.8893 | -104.5061
8.893 | -104.5166
8.8966 | -104.5236
8.9003 | -104.5318
8.904 | -104.5478
8.9077 | -104.5674
8.9113 | -104.5867
8.915 | -104.6039
8.9187 | -104.6245
8.9224 | -104.6477
8.926 | -104.6701
8.9297 | -104.6953
8.9334 | -104.7178
8.937 | -104.7396
8.9407 | -104.7623
8.9444 | -104.7869
8.9481 | -104.8084
8.9517 | -104.8278
8.9554 | -104.8394
8.9591 | -104.8479
8.9628 | -104.8578
8.9664 | -104.861
8.9701 | -104.8636
8.9738 | -104.8654
8.9775 | -104.8629
8.9811 | -104.8486
8.9848 | -104.8349
8.9885 | -104.8159
8.9921 | -104.7895
8.9958 | -104.7628
8.9995 | -104.7357
9.0032 | -104.7095
9.0068 | -104.6904
9.0105 | -104.6712
9.0142 | -104.6561
9.0179 | -104.6464
9.0215 | -104.6403
9.0252 | -104.6374
9.0289 | -104.6429
9.0326 | -104.6478
9.0362 | -104.6514
9.0399 | -104.6616
9.0436 | -104.6774
9.0472 | -104.695
9.0509 | -104.7124
9.0546 | -104.7312
9.0583 | -104.7499
9.0619 | -104.7593
9.0656 | -104.7681
9.0693 | -104.7781
9.073 | -104.7808
9.0766 | -104.7769
9.0803 | -104.7735
9.084 | -104.7632
9.0877 | -104.7451
9.0913 | -104.7221
9.095 | -104.6921
9.0987 | -104.656
9.1023 | -104.6144
9.106 | -104.573
9.1097 | -104.5316
9.1134 | -104.4963
9.117 | -104.462
9.1207 | -104.433
9.1244 | -104.408
9.1281 | -104.3866
9.1317 | -104.3693
9.1354 | -104.3536
9.1391 | -104.3449
9.1428 | -104.3421
9.1464 | -104.3409
9.1501 | -104.3433
9.1538 | -104.3504
9.1574 | -104.3607
9.1611 | -104.3731
9.1648 | -104.3877
9.1685 | -104.4095
9.1721 | -104.434
9.1758 | -104.4555
9.1795 | -104.4833
9.1832 | -104.5148
9.1868 | -104.5481
9.1905 | -104.5828
9.1942 | -104.6131
9.1979 | -104.6459
9.2015 | -104.6754
9.2052 | -104.7
9.2089 | -104.7201
9.2125 | -104.7348
9.2162 | -104.7458
9.2199 | -104.7529
9.2236 | -104.762
9.2272 | -104.7667
9.2309 | -104.7687
9.2346 | -104.7738
9.2383 | -104.7719
9.2419 | -104.7621
9.2456 | -104.7509
9.2493 | -104.7372
9.253 | -104.7208
9.2566 | -104.7003
9.2603 | -104.6743
9.264 | -104.649
9.2676 | -104.6249
9.2713 | -104.5992
9.275 | -104.5693
9.2787 | -104.5376
9.2823 | -104.5073
9.286 | -104.4757
9.2897 | -104.4401
9.2934 | -104.4053
9.297 | -104.3725
9.3007 | -104.3415
9.3044 | -104.3093
9.308 | -104.2795
9.3117 | -104.2577
9.3154 | -104.2365
9.3191 | -104.2216
9.3227 | -104.2122
9.3264 | -104.2086
9.3301 | -104.2136
9.3338 | -104.2219
9.3374 | -104.2392
9.3411 | -104.2604
9.3448 | -104.2858
9.3485 | -104.3156
9.3521 | -104.3491
9.3558 | -104.3849
9.3595 | -104.42
9.3631 | -104.4532
9.3668 | -104.4895
9.3705 | -104.5262
9.3742 | -104.5665
9.3778 | -104.6038
9.3815 | -104.6399
9.3852 | -104.6733
9.3889 | -104.7038
9.3925 | -104.7333
9.3962 | -104.7647
9.3999 | -104.7853
9.4036 | -104.8039
9.4072 | -104.8194
9.4109 | -104.8304
9.4146 | -104.8367
9.4182 | -104.8348
9.4219 | -104.8309
9.4256 | -104.828
9.4293 | -104.8209
9.4329 | -104.8091
9.4366 | -104.7934
9.4403 | -104.7805
9.444 | -104.7684
9.4476 | -104.7592
9.4513 | -104.7522
9.455 | -104.7455
9.4587 | -104.7409
9.4623 | -104.7381
9.466 | -104.7354
9.4697 | -104.736
9.4733 | -104.734
9.477 | -104.735
9.4807 | -104.7358
9.4844 | -104.7361
9.488 | -104.7295
9.4917 | -104.7237
9.4954 | -104.7177
9.4991 | -104.7126
9.5027 | -104.7058
9.5064 | -104.6951
9.5101 | -104.6882
9.5138 | -104.685
9.5174 | -104.6839
9.5211 | -104.6862
9.5248 | -104.6891
9.5284 | -104.6965
9.5321 | -104.7098
9.5358 | -104.7341
9.5395 | -104.7546
9.5431 | -104.7717
9.5468 | -104.7933
9.5505 | -104.8198
9.5542 | -104.8439
9.5578 | -104.8676
9.5615 | -104.8896
9.5652 | -104.904
9.5689 | -104.9218
9.5725 | -104.935
9.5762 | -104.9428
9.5799 | -104.9481
9.5835 | -104.9513
9.5872 | -104.9453
9.5909 | -104.94
9.5946 | -104.9334
9.5982 | -104.9239
9.6019 | -104.913
9.6056 | -104.9021
9.6093 | -104.8873
9.6129 | -104.871
9.6166 | -104.8539
9.6203 | -104.8396
9.624 | -104.8202
9.6276 | -104.8018
9.6313 | -104.7811
9.635 | -104.7621
9.6386 | -104.7397
9.6423 | -104.7183
9.646 | -104.7006
9.6497 | -104.6778
9.6533 | -104.6598
9.657 | -104.6425
9.6607 | -104.6262
9.6644 | -104.6191
9.668 | -104.6153
9.6717 | -104.6143
9.6754 | -104.6195
9.679 | -104.6277
9.6827 | -104.6382
9.6864 | -104.6537
9.6901 | -104.6724
9.6937 | -104.692
9.6974 | -104.7177
9.7011 | -104.7459
9.7048 | -104.7738
9.7084 | -104.8059
9.7121 | -104.8371
9.7158 | -104.8672
9.7195 | -104.8961
9.7231 | -104.9248
9.7268 | -104.9541
9.7305 | -104.9805
9.7341 | -105.0001
9.7378 | -105.0145
9.7415 | -105.0255
9.7452 | -105.029
9.7488 | -105.0226
9.7525 | -105.0144
9.7562 | -105.0054
9.7599 | -104.9913
9.7635 | -104.9703
9.7672 | -104.9517
9.7709 | -104.9336
9.7746 | -104.9166
9.7782 | -104.8988
9.7819 | -104.8891
9.7856 | -104.8837
9.7892 | -104.8782
9.7929 | -104.8788
9.7966 | -104.8782
9.8003 | -104.881
9.8039 | -104.8874
9.8076 | -104.8962
9.8113 | -104.9083
9.815 | -104.9177
9.8186 | -104.9286
9.8223 | -104.9392
9.826 | -104.9525
9.8297 | -104.9652
9.8333 | -104.974
9.837 | -104.9818
9.8407 | -104.9859
9.8443 | -104.9889
9.848 | -104.9915
9.8517 | -104.9926
9.8554 | -104.9932
9.859 | -104.9956
9.8627 | -104.9949
9.8664 | -104.988
9.8701 | -104.9814
9.8737 | -104.9767
9.8774 | -104.973
9.8811 | -104.9742
9.8848 | -104.9752
9.8884 | -104.98
9.8921 | -104.9897
9.8958 | -104.9961
9.8994 | -105.0018
9.9031 | -105.015
9.9068 | -105.0365
9.9105 | -105.0568
9.9141 | -105.0801
9.9178 | -105.1025
9.9215 | -105.1302
9.9252 | -105.1552
9.9288 | -105.1792
9.9325 | -105.2073
9.9362 | -105.2308
9.9398 | -105.249
9.9435 | -105.2628
9.9472 | -105.2717
9.9509 | -105.2761
9.9545 | -105.2788
9.9582 | -105.2836
9.9619 | -105.2836
9.9656 | -105.2783
9.9692 | -105.2706
9.9729 | -105.2676
9.9766 | -105.2688
9.9803 | -105.269
9.9839 | -105.2677
9.9876 | -105.2651
9.9913 | -105.2646
9.9949 | -105.2663
9.9986 | -105.2683
10.0023 | -105.2743
10.006 | -105.2829
10.0096 | -105.2922
10.0133 | -105.2947
10.017 | -105.2973
10.0207 | -105.2964
10.0243 | -105.291
10.028 | -105.2884
10.0317 | -105.2827
10.0354 | -105.2746
10.039 | -105.267
10.0427 | -105.2559
10.0464 | -105.24
10.05 | -105.2231
10.0537 | -105.2102
10.0574 | -105.196
10.0611 | -105.1859
10.0647 | -105.1721
10.0684 | -105.1682
10.0721 | -105.1605
10.0758 | -105.1542
10.0794 | -105.1456
10.0831 | -105.1365
10.0868 | -105.131
10.0905 | -105.1255
10.0941 | -105.1184
10.0978 | -105.105
10.1015 | -105.086
10.1051 | -105.0702
10.1088 | -105.0574
10.1125 | -105.0434
10.1162 | -105.0243
10.1198 | -105.003
10.1235 | -104.9792
10.1272 | -104.9587
10.1309 | -104.9395
10.1345 | -104.9202
10.1382 | -104.8975
10.1419 | -104.8752
10.1456 | -104.854
10.1492 | -104.8341
10.1529 | -104.8133
10.1566 | -104.7963
10.1602 | -104.7858
10.1639 | -104.7744
10.1676 | -104.7643
10.1713 | -104.7569
10.1749 | -104.7516
10.1786 | -104.75
10.1823 | -104.747
10.186 | -104.7449
10.1896 | -104.7456
10.1933 | -104.7486
10.197 | -104.7516
10.2007 | -104.7558
10.2043 | -104.7614
10.208 | -104.7666
10.2117 | -104.7717
10.2153 | -104.7752
10.219 | -104.7834
10.2227 | -104.7934
10.2264 | -104.8073
10.23 | -104.8246
10.2337 | -104.8467
10.2374 | -104.8707
10.2411 | -104.8959
10.2447 | -104.9229
10.2484 | -104.9526
10.2521 | -104.9871
10.2558 | -105.0195
10.2594 | -105.0515
10.2631 | -105.0882
10.2668 | -105.1275
10.2704 | -105.1608
10.2741 | -105.194
10.2778 | -105.2277
10.2815 | -105.261
10.2851 | -105.2924
10.2888 | -105.3199
10.2925 | -105.3439
10.2962 | -105.3613
10.2998 | -105.3734
10.3035 | -105.3881
10.3072 | -105.3945
10.3109 | -105.3951
10.3145 | -105.3929
10.3182 | -105.3884
10.3219 | -105.3775
10.3255 | -105.364
10.3292 | -105.3472
10.3329 | -105.3343
10.3366 | -105.3258
10.3402 | -105.3128
10.3439 | -105.3001
10.3476 | -105.2895
10.3513 | -105.2802
10.3549 | -105.2747
10.3586 | -105.2723
10.3623 | -105.2735
10.3659 | -105.2774
10.3696 | -105.2791
10.3733 | -105.2824
10.377 | -105.2876
10.3806 | -105.2954
10.3843 | -105.2992
10.388 | -105.308
10.3917 | -105.3225
10.3953 | -105.3386
10.399 | -105.3543
10.4027 | -105.3673
10.4064 | -105.3806
10.41 | -105.3955
10.4137 | -105.4036
10.4174 | -105.4148
10.421 | -105.4304
10.4247 | -105.443
10.4284 | -105.4505
10.4321 | -105.4541
10.4357 | -105.4514
10.4394 | -105.452
10.4431 | -105.4529
10.4468 | -105.4503
10.4504 | -105.4461
10.4541 | -105.4355
10.4578 | -105.4219
10.4615 | -105.4014
10.4651 | -105.3846
10.4688 | -105.366
10.4725 | -105.3447
10.4761 | -105.3243
10.4798 | -105.3051
10.4835 | -105.2855
10.4872 | -105.2659
10.4908 | -105.2519
10.4945 | -105.2388
10.4982 | -105.226
10.5019 | -105.2162
10.5055 | -105.208
10.5092 | -105.2036
10.5129 | -105.1976
10.5166 | -105.1931
10.5202 | -105.1912
10.5239 | -105.1903
10.5276 | -105.1838
10.5312 | -105.1796
10.5349 | -105.1759
10.5386 | -105.1675
10.5423 | -105.1556
10.5459 | -105.1455
10.5496 | -105.1348
10.5533 | -105.1201
10.557 | -105.106
10.5606 | -105.091
10.5643 | -105.0768
10.568 | -105.0611
10.5717 | -105.0446
10.5753 | -105.0282
10.579 | -105.016
10.5827 | -105.0075
10.5863 | -105.0024
10.59 | -104.9971
10.5937 | -104.9919
10.5974 | -105.0054
10.601 | -105.0079
10.6047 | -105.0127
10.6084 | -105.0211
10.6121 | -105.0342
10.6157 | -105.048
10.6194 | -105.0588
10.6231 | -105.0707
10.6268 | -105.088
10.6304 | -105.103
10.6341 | -105.1182
10.6378 | -105.1375
10.6414 | -105.1555
10.6451 | -105.1737
10.6488 | -105.1905
10.6525 | -105.2085
10.6561 | -105.2281
10.6598 | -105.2497
10.6635 | -105.2695
10.6672 | -105.2964
10.6708 | -105.3243
10.6745 | -105.3495
10.6782 | -105.3755
10.6818 | -105.4061
10.6855 | -105.432
10.6892 | -105.4533
10.6929 | -105.4693
10.6965 | -105.4868
10.7002 | -105.5031
10.7039 | -105.5151
10.7076 | -105.525
10.7112 | -105.5317
10.7149 | -105.5315
10.7186 | -105.5321
10.7223 | -105.5288
10.7259 | -105.5224
10.7296 | -105.5141
10.7333 | -105.5034
10.7369 | -105.4898
10.7406 | -105.475
10.7443 | -105.4575
10.748 | -105.4364
10.7516 | -105.4157
10.7553 | -105.3981
10.759 | -105.3797
10.7627 | -105.357
10.7663 | -105.3309
10.77 | -105.3072
10.7737 | -105.2863
10.7774 | -105.2655
10.781 | -105.2467
10.7847 | -105.2258
10.7884 | -105.2086
10.792 | -105.1916
10.7957 | -105.1702
10.7994 | -105.1555
10.8031 | -105.1453
10.8067 | -105.1354
10.8104 | -105.1303
10.8141 | -105.1259
10.8178 | -105.1223
10.8214 | -105.1204
10.8251 | -105.1154
10.8288 | -105.1156
10.8325 | -105.1218
10.8361 | -105.1272
10.8398 | -105.1361
10.8435 | -105.1448
10.8471 | -105.1592
10.8508 | -105.1757
10.8545 | -105.187
10.8582 | -105.1985
10.8618 | -105.2137
10.8655 | -105.2302
10.8692 | -105.2454
10.8729 | -105.267
10.8765 | -105.2868
10.8802 | -105.3016
10.8839 | -105.3188
10.8876 | -105.337
10.8912 | -105.3521
10.8949 | -105.3647
10.8986 | -105.3773
10.9022 | -105.3895
10.9059 | -105.4026
10.9096 | -105.4128
10.9133 | -105.4179
10.9169 | -105.4225
10.9206 | -105.4297
10.9243 | -105.4328
10.928 | -105.4318
10.9316 | -105.4282
10.9353 | -105.4233
10.939 | -105.4173
10.9427 | -105.4125
10.9463 | -105.4039
10.95 | -105.398
10.9537 | -105.3951
10.9573 | -105.3868
10.961 | -105.3757
10.9647 | -105.3653
10.9684 | -105.3541
10.972 | -105.3429
10.9757 | -105.332
10.9794 | -105.3225
10.9831 | -105.3153
10.9867 | -105.3076
10.9904 | -105.2947
10.9941 | -105.286
10.9977 | -105.2789
11.0014 | -105.2782
11.0051 | -105.275
11.0088 | -105.2694
11.0124 | -105.2667
11.0161 | -105.2639
11.0198 | -105.2576
11.0235 | -105.2502
11.0271 | -105.2449
11.0308 | -105.235
11.0345 | -105.2267
11.0382 | -105.2246
11.0418 | -105.2219
11.0455 | -105.217
11.0492 | -105.208
11.0528 | -105.2037
11.0565 | -105.1985
11.0602 | -105.1943
11.0639 | -105.1932
11.0675 | -105.1917
11.0712 | -105.1942
11.0749 | -105.1975
11.0786 | -105.2033
11.0822 | -105.2124
11.0859 | -105.2199
11.0896 | -105.2314
11.0933 | -105.2448
11.0969 | -105.2587
11.1006 | -105.2708
11.1043 | -105.285
11.1079 | -105.297
11.1116 | -105.3079
11.1153 | -105.3186
11.119 | -105.3334
11.1226 | -105.3528
11.1263 | -105.3604
11.13 | -105.367
11.1337 | -105.3716
11.1373 | -105.376
11.141 | -105.3806
11.1447 | -105.3824
11.1484 | -105.3831
11.152 | -105.3864
11.1557 | -105.3843
11.1594 | -105.3812
11.163 | -105.3765
11.1667 | -105.3758
11.1704 | -105.3746
11.1741 | -105.37
11.1777 | -105.363
11.1814 | -105.3628
11.1851 | -105.3627
11.1888 | -105.3592
11.1924 | -105.3604
11.1961 | -105.3625
11.1998 | -105.3652
11.2035 | -105.3701
11.2071 | -105.3788
11.2108 | -105.3924
11.2145 | -105.4042
11.2181 | -105.4154
11.2218 | -105.4249
11.2255 | -105.4365
11.2292 | -105.4501
11.2328 | -105.4612
11.2365 | -105.4723
11.2402 | -105.4789
11.2439 | -105.4847
11.2475 | -105.4914
11.2512 | -105.4954
11.2549 | -105.4937
11.2586 | -105.489
11.2622 | -105.4811
11.2659 | -105.4703
11.2696 | -105.4607
11.2732 | -105.4478
11.2769 | -105.4331
11.2806 | -105.4102
11.2843 | -105.3847
11.2879 | -105.364
11.2916 | -105.3456
11.2953 | -105.3279
11.299 | -105.308
11.3026 | -105.2886
11.3063 | -105.2674
11.31 | -105.2462
11.3137 | -105.2257
11.3173 | -105.2082
11.321 | -105.1938
11.3247 | -105.1793
11.3283 | -105.1667
11.332 | -105.1571
11.3357 | -105.1497
11.3394 | -105.1449
11.343 | -105.1425
11.3467 | -105.1388
11.3504 | -105.1394
11.3541 | -105.1442
11.3577 | -105.1517
11.3614 | -105.1618
11.3651 | -105.1719
11.3688 | -105.1849
11.3724 | -105.2009
11.3761 | -105.2223
11.3798 | -105.2434
11.3834 | -105.2699
11.3871 | -105.298
11.3908 | -105.3249
11.3945 | -105.3512
11.3981 | -105.3837
11.4018 | -105.4134
11.4055 | -105.4431
11.4092 | -105.4708
11.4128 | -105.4932
11.4165 | -105.5116
11.4202 | -105.5313
11.4238 | -105.5479
11.4275 | -105.5639
11.4312 | -105.5755
11.4349 | -105.5801
11.4385 | -105.5779
11.4422 | -105.5742
11.4459 | -105.5722
11.4496 | -105.5643
11.4532 | -105.5533
11.4569 | -105.5371
11.4606 | -105.5198
11.4643 | -105.5032
11.4679 | -105.4819
11.4716 | -105.4603
11.4753 | -105.4385
11.4789 | -105.4172
11.4826 | -105.3955
11.4863 | -105.373
11.49 | -105.3547
11.4936 | -105.3402
11.4973 | -105.3262
11.501 | -105.3144
11.5047 | -105.3094
11.5083 | -105.3066
11.512 | -105.3042
11.5157 | -105.3044
11.5194 | -105.3074
11.523 | -105.3139
11.5267 | -105.3222
11.5304 | -105.3347
11.534 | -105.3456
11.5377 | -105.3592
11.5414 | -105.3725
11.5451 | -105.3877
11.5487 | -105.4033
11.5524 | -105.4185
11.5561 | -105.4358
11.5598 | -105.4571
11.5634 | -105.4751
11.5671 | -105.4957
11.5708 | -105.5144
11.5745 | -105.5337
11.5781 | -105.555
11.5818 | -105.5753
11.5855 | -105.5917
11.5891 | -105.6103
11.5928 | -105.628
11.5965 | -105.6449
11.6002 | -105.6619
11.6038 | -105.6779
11.6075 | -105.6856
11.6112 | -105.6924
11.6149 | -105.7004
11.6185 | -105.7049
11.6222 | -105.7092
11.6259 | -105.7149
11.6296 | -105.7216
11.6332 | -105.7262
11.6369 | -105.73
11.6406 | -105.7321
11.6442 | -105.7328
11.6479 | -105.7367
11.6516 | -105.7409
11.6553 | -105.7386
11.6589 | -105.738
11.6626 | -105.7349
11.6663 | -105.7324
11.67 | -105.7335
11.6736 | -105.7322
11.6773 | -105.7325
11.681 | -105.7357
11.6846 | -105.7333
11.6883 | -105.7321
11.692 | -105.7323
11.6957 | -105.7369
11.6993 | -105.745
11.703 | -105.7544
11.7067 | -105.7625
11.7104 | -105.7671
11.714 | -105.773
11.7177 | -105.7797
11.7214 | -105.7899
11.7251 | -105.8018
11.7287 | -105.809
11.7324 | -105.8141
11.7361 | -105.8194
11.7397 | -105.827
11.7434 | -105.8352
11.7471 | -105.8421
11.7508 | -105.8454
11.7544 | -105.8453
11.7581 | -105.8456
11.7618 | -105.8418
11.7655 | -105.8369
11.7691 | -105.8343
11.7728 | -105.8337
11.7765 | -105.8335
11.7802 | -105.8279
11.7838 | -105.8222
11.7875 | -105.8171
11.7912 | -105.811
11.7948 | -105.8059
11.7985 | -105.8006
11.8022 | -105.8002
11.8059 | -105.8011
11.8095 | -105.803
11.8132 | -105.8045
11.8169 | -105.8051
11.8206 | -105.8135
11.8242 | -105.8265
11.8279 | -105.8411
11.8316 | -105.856
11.8353 | -105.8744
11.8389 | -105.8922
11.8426 | -105.9074
11.8463 | -105.9233
11.8499 | -105.9385
11.8536 | -105.9569
11.8573 | -105.9793
11.861 | -105.9936
11.8646 | -106.006
11.8683 | -106.0137
11.872 | -106.0149
11.8757 | -106.017
11.8793 | -106.0203
11.883 | -106.0255
11.8867 | -106.0278
11.8904 | -106.0237
11.894 | -106.0138
11.8977 | -106.0007
11.9014 | -105.9876
11.905 | -105.9733
11.9087 | -105.9593
11.9124 | -105.9466
11.9161 | -105.9301
11.9197 | -105.9156
11.9234 | -105.9055
11.9271 | -105.8941
11.9308 | -105.88
11.9344 | -105.873
11.9381 | -105.87
11.9418 | -105.8647
11.9455 | -105.8612
11.9491 | -105.863
11.9528 | -105.8682
11.9565 | -105.8707
11.9601 | -105.8756
11.9638 | -105.878
11.9675 | -105.8787
11.9712 | -105.8784
11.9748 | -105.8795
11.9785 | -105.8812
11.9822 | -105.8832
11.9859 | -105.8797
11.9895 | -105.8755
11.9932 | -105.8652
11.9969 | -105.8543
12.0006 | -105.8431
12.0042 | -105.8299
12.0079 | -105.8159
12.0116 | -105.8025
12.0152 | -105.791
12.0189 | -105.7761
12.0226 | -105.7599
12.0263 | -105.7387
12.0299 | -105.7206
12.0336 | -105.71
12.0373 | -105.6989
12.041 | -105.6911
12.0446 | -105.6862
12.0483 | -105.6818
12.052 | -105.6767
12.0556 | -105.6769
12.0593 | -105.6842
12.063 | -105.6952
12.0667 | -105.7121
12.0703 | -105.7319
12.074 | -105.755
12.0777 | -105.782
12.0814 | -105.8063
12.085 | -105.8281
12.0887 | -105.8548
12.0924 | -105.8811
12.0961 | -105.906
12.0997 | -105.9338
12.1034 | -105.958
12.1071 | -105.9752
12.1107 | -105.9884
12.1144 | -106.0032
12.1181 | -106.0159
12.1218 | -106.0223
12.1254 | -106.0276
12.1291 | -106.031
12.1328 | -106.0314
12.1365 | -106.032
12.1401 | -106.0298
12.1438 | -106.0265
12.1475 | -106.0196
12.1512 | -106.0092
12.1548 | -105.9975
12.1585 | -105.9909
12.1622 | -105.985
12.1658 | -105.9805
12.1695 | -105.9725
12.1732 | -105.965
12.1769 | -105.9609
12.1805 | -105.9537
12.1842 | -105.9467
12.1879 | -105.9449
12.1916 | -105.9373
12.1952 | -105.9348
12.1989 | -105.933
12.2026 | -105.9327
12.2063 | -105.932
12.2099 | -105.9299
12.2136 | -105.9306
12.2173 | -105.9315
12.2209 | -105.932
12.2246 | -105.9322
12.2283 | -105.9351
12.232 | -105.9361
12.2356 | -105.9335
12.2393 | -105.9308
12.243 | -105.9322
12.2467 | -105.9321
12.2503 | -105.9331
12.254 | -105.9369
12.2577 | -105.9426
12.2614 | -105.9497
12.265 | -105.953
12.2687 | -105.9567
12.2724 | -105.962
12.276 | -105.9664
12.2797 | -105.9729
12.2834 | -105.9811
12.2871 | -105.9881
12.2907 | -105.9924
12.2944 | -105.9994
12.2981 | -106.0072
12.3018 | -106.0106
12.3054 | -106.0137
12.3091 | -106.014
12.3128 | -106.0164
12.3165 | -106.0189
12.3201 | -106.0152
12.3238 | -106.0076
12.3275 | -106.0032
12.3311 | -105.9975
12.3348 | -105.9868
12.3385 | -105.9739
12.3422 | -105.9617
12.3458 | -105.9496
12.3495 | -105.937
12.3532 | -105.9216
12.3569 | -105.9077
12.3605 | -105.8921
12.3642 | -105.8736
12.3679 | -105.8577
12.3716 | -105.8426
12.3752 | -105.824
12.3789 | -105.8053
12.3826 | -105.786
12.3862 | -105.7674
12.3899 | -105.752
12.3936 | -105.7338
12.3973 | -105.7146
12.4009 | -105.7019
12.4046 | -105.6902
12.4083 | -105.6773
12.412 | -105.6646
12.4156 | -105.6533
12.4193 | -105.6432
12.423 | -105.6319
12.4267 | -105.6233
12.4303 | -105.6158
12.434 | -105.6061
12.4377 | -105.5963
12.4413 | -105.5874
12.445 | -105.5799
12.4487 | -105.5728
12.4524 | -105.568
12.456 | -105.5652
12.4597 | -105.5621
12.4634 | -105.5609
12.4671 | -105.562
12.4707 | -105.5679
12.4744 | -105.574
12.4781 | -105.582
12.4817 | -105.5937
12.4854 | -105.6104
12.4891 | -105.6311
12.4928 | -105.6511
12.4964 | -105.6737
12.5001 | -105.7002
12.5038 | -105.73
12.5075 | -105.7603
12.5111 | -105.792
12.5148 | -105.8276
12.5185 | -105.8606
12.5222 | -105.8931
12.5258 | -105.9297
12.5295 | -105.964
12.5332 | -105.9979
12.5368 | -106.0303
12.5405 | -106.0603
12.5442 | -106.0886
12.5479 | -106.1119
12.5515 | -106.1308
12.5552 | -106.1473
12.5589 | -106.1631
12.5626 | -106.1742
12.5662 | -106.1772
12.5699 | -106.1754
12.5736 | -106.1778
12.5773 | -106.179
12.5809 | -106.1731
12.5846 | -106.1644
12.5883 | -106.1528
12.5919 | -106.1434
12.5956 | -106.1325
12.5993 | -106.1216
12.603 | -106.1093
12.6066 | -106.0962
12.6103 | -106.0853
12.614 | -106.0766
12.6177 | -106.0705
12.6213 | -106.0629
12.625 | -106.0563
12.6287 | -106.0452
12.6324 | -106.0352
12.636 | -106.0257
12.6397 | -106.0178
12.6434 | -106.0098
12.647 | -106.0028
12.6507 | -105.9964
12.6544 | -105.9922
12.6581 | -105.989
12.6617 | -105.9838
12.6654 | -105.9763
12.6691 | -105.9693
12.6728 | -105.963
12.6764 | -105.9558
12.6801 | -105.9514
12.6838 | -105.9461
12.6875 | -105.9378
12.6911 | -105.9349
12.6948 | -105.9326
12.6985 | -105.9299
12.7021 | -105.9262
12.7058 | -105.9243
12.7095 | -105.9251
12.7132 | -105.9273
12.7168 | -105.9371
12.7205 | -105.9425
12.7242 | -105.9539
12.7279 | -105.9624
12.7315 | -105.9726
12.7352 | -105.9825
12.7389 | -105.9943
12.7425 | -106.0048
12.7462 | -106.0148
12.7499 | -106.03
12.7536 | -106.0415
12.7572 | -106.0489
12.7609 | -106.0576
12.7646 | -106.0693
12.7683 | -106.0816
12.7719 | -106.0919
12.7756 | -106.1005
12.7793 | -106.1115
12.783 | -106.1208
12.7866 | -106.1276
12.7903 | -106.1353
12.794 | -106.1434
12.7976 | -106.1501
12.8013 | -106.1566
12.805 | -106.1651
12.8087 | -106.1742
12.8123 | -106.1835
12.816 | -106.1878
12.8197 | -106.1934
12.8234 | -106.1955
12.827 | -106.1984
12.8307 | -106.2019
12.8344 | -106.2037
12.8381 | -106.2041
12.8417 | -106.2049
12.8454 | -106.2089
12.8491 | -106.2109
12.8527 | -106.2119
12.8564 | -106.213
12.8601 | -106.2152
12.8638 | -106.2179
12.8674 | -106.2218
12.8711 | -106.225
12.8748 | -106.2227
12.8785 | -106.2217
12.8821 | -106.2238
12.8858 | -106.2278
12.8895 | -106.23
12.8932 | -106.2314
12.8968 | -106.2328
12.9005 | -106.2328
12.9042 | -106.2313
12.9078 | -106.2324
12.9115 | -106.2277
12.9152 | -106.2261
12.9189 | -106.2274
12.9225 | -106.23
12.9262 | -106.2284
12.9299 | -106.2265
12.9336 | -106.2288
12.9372 | -106.2252
12.9409 | -106.22
12.9446 | -106.2136
12.9483 | -106.207
12.9519 | -106.2
12.9556 | -106.1911
12.9593 | -106.1857
12.9629 | -106.1775
12.9666 | -106.167
12.9703 | -106.1544
12.974 | -106.1418
12.9776 | -106.1269
12.9813 | -106.1099
12.985 | -106.0949
12.9887 | -106.0803
12.9923 | -106.066
12.996 | -106.0503
12.9997 | -106.0384
13.0034 | -106.0273
13.007 | -106.0138
13.0107 | -106.0051
13.0144 | -105.9955
13.018 | -105.9869
13.0217 | -105.9771
13.0254 | -105.9712
13.0291 | -105.968
13.0327 | -105.9633
13.0364 | -105.9656
13.0401 | -105.9702
13.0438 | -105.9778
13.0474 | -105.9873
13.0511 | -105.9969
13.0548 | -106.0036
13.0585 | -106.0098
13.0621 | -106.0185
13.0658 | -106.0272
13.0695 | -106.0376
13.0731 | -106.0447
13.0768 | -106.0471
13.0805 | -106.0517
13.0842 | -106.0534
13.0878 | -106.0525
13.0915 | -106.0487
13.0952 | -106.0442
13.0989 | -106.0414
13.1025 | -106.0364
13.1062 | -106.0272
13.1099 | -106.0174
13.1136 | -106.0087
13.1172 | -105.9998
13.1209 | -105.9848
13.1246 | -105.9715
13.1282 | -105.9605
13.1319 | -105.9479
13.1356 | -105.9349
13.1393 | -105.9207
13.1429 | -105.9058
13.1466 | -105.8915
13.1503 | -105.8802
13.154 | -105.8691
13.1576 | -105.8605
13.1613 | -105.8557
13.165 | -105.8511
13.1686 | -105.8457
13.1723 | -105.8417
13.176 | -105.8393
13.1797 | -105.8378
13.1833 | -105.8403
13.187 | -105.8435
13.1907 | -105.8463
13.1944 | -105.8518
13.198 | -105.8546
13.2017 | -105.8622
13.2054 | -105.8699
13.2091 | -105.8782
13.2127 | -105.8861
13.2164 | -105.8925
13.2201 | -105.899
13.2237 | -105.9003
13.2274 | -105.8989
13.2311 | -105.8948
13.2348 | -105.8905
13.2384 | -105.8854
13.2421 | -105.8831
13.2458 | -105.8776
13.2495 | -105.8791
13.2531 | -105.8767
13.2568 | -105.8719
13.2605 | -105.8679
13.2642 | -105.8656
13.2678 | -105.8665
13.2715 | -105.8652
13.2752 | -105.8666
13.2788 | -105.8702
13.2825 | -105.8748
13.2862 | -105.8785
13.2899 | -105.8819
13.2935 | -105.8897
13.2972 | -105.8989
13.3009 | -105.9087
13.3046 | -105.9225
13.3082 | -105.9379
13.3119 | -105.9586
13.3156 | -105.9797
13.3193 | -106.0012
13.3229 | -106.0256
13.3266 | -106.0492
13.3303 | -106.0708
13.3339 | -106.0944
13.3376 | -106.117
13.3413 | -106.1387
13.345 | -106.1608
13.3486 | -106.1811
13.3523 | -106.1993
13.356 | -106.214
13.3597 | -106.229
13.3633 | -106.2413
13.367 | -106.2516
13.3707 | -106.2605
13.3744 | -106.2656
13.378 | -106.2686
13.3817 | -106.2693
13.3854 | -106.2641
13.389 | -106.2599
13.3927 | -106.2528
13.3964 | -106.2425
13.4001 | -106.2341
13.4037 | -106.223
13.4074 | -106.2126
13.4111 | -106.2044
13.4148 | -106.1957
13.4184 | -106.1872
13.4221 | -106.1769
13.4258 | -106.1669
13.4294 | -106.1591
13.4331 | -106.1526
13.4368 | -106.149
13.4405 | -106.1462
13.4441 | -106.1453
13.4478 | -106.147
13.4515 | -106.1518
13.4552 | -106.1592
13.4588 | -106.1677
13.4625 | -106.1771
13.4662 | -106.1879
13.4699 | -106.2011
13.4735 | -106.2154
13.4772 | -106.2289
13.4809 | -106.2443
13.4845 | -106.2617
13.4882 | -106.2758
13.4919 | -106.2908
13.4956 | -106.3076
13.4992 | -106.32
13.5029 | -106.3283
13.5066 | -106.3381
13.5103 | -106.3459
13.5139 | -106.3504
13.5176 | -106.3533
13.5213 | -106.3552
13.525 | -106.3559
13.5286 | -106.3586
13.5323 | -106.3564
13.536 | -106.3499
13.5396 | -106.3413
13.5433 | -106.3336
13.547 | -106.3199
13.5507 | -106.3092
13.5543 | -106.2966
13.558 | -106.2818
13.5617 | -106.2636
13.5654 | -106.2436
13.569 | -106.2242
13.5727 | -106.2074
13.5764 | -106.1932
13.5801 | -106.1791
13.5837 | -106.1602
13.5874 | -106.1432
13.5911 | -106.1296
13.5947 | -106.1164
13.5984 | -106.1031
13.6021 | -106.0921
13.6058 | -106.0828
13.6094 | -106.0773
13.6131 | -106.073
13.6168 | -106.0702
13.6205 | -106.0697
13.6241 | -106.07
13.6278 | -106.0705
13.6315 | -106.0694
13.6352 | -106.0706
13.6388 | -106.0749
13.6425 | -106.0807
13.6462 | -106.0882
13.6498 | -106.0946
13.6535 | -106.1019
13.6572 | -106.1077
13.6609 | -106.1135
13.6645 | -106.1195
13.6682 | -106.1239
13.6719 | -106.1295
13.6756 | -106.1318
13.6792 | -106.134
13.6829 | -106.1373
13.6866 | -106.1375
13.6903 | -106.1344
13.6939 | -106.1335
13.6976 | -106.1344
13.7013 | -106.131
13.7049 | -106.1274
13.7086 | -106.1229
13.7123 | -106.1146
13.716 | -106.106
13.7196 | -106.0999
13.7233 | -106.0883
13.727 | -106.0777
13.7307 | -106.0679
13.7343 | -106.0561
13.738 | -106.0452
13.7417 | -106.0365
13.7454 | -106.0263
13.749 | -106.0143
13.7527 | -106.0028
13.7564 | -105.9931
13.76 | -105.986
13.7637 | -105.9777
13.7674 | -105.9663
13.7711 | -105.9565
13.7747 | -105.9508
13.7784 | -105.9438
13.7821 | -105.9404
13.7858 | -105.9337
13.7894 | -105.9292
13.7931 | -105.9253
13.7968 | -105.9233
13.8004 | -105.9188
13.8041 | -105.9146
13.8078 | -105.9125
13.8115 | -105.9141
13.8151 | -105.9183
13.8188 | -105.9203
13.8225 | -105.9207
13.8262 | -105.922
13.8298 | -105.9268
13.8335 | -105.9335
13.8372 | -105.9401
13.8409 | -105.9455
13.8445 | -105.9509
13.8482 | -105.957
13.8519 | -105.9634
13.8555 | -105.9742
13.8592 | -105.985
13.8629 | -105.9902
13.8666 | -105.9939
13.8702 | -105.9999
13.8739 | -106.0068
13.8776 | -106.013
13.8813 | -106.0173
13.8849 | -106.0215
13.8886 | -106.0243
13.8923 | -106.0238
13.896 | -106.0219
13.8996 | -106.0225
13.9033 | -106.0229
13.907 | -106.0249
13.9106 | -106.0272
13.9143 | -106.0272
13.918 | -106.0251
13.9217 | -106.0251
13.9253 | -106.0235
13.929 | -106.0199
13.9327 | -106.0198
13.9364 | -106.016
13.94 | -106.0124
13.9437 | -106.0106
13.9474 | -106.0111
13.9511 | -106.0114
13.9547 | -106.0097
13.9584 | -106.0001
13.9621 | -105.9848
13.9657 | -105.9711
13.9694 | -105.9599
13.9731 | -105.9509
13.9768 | -105.9421
13.9804 | -105.932
13.9841 | -105.9188
13.9878 | -105.905
13.9915 | -105.8881
13.9951 | -105.8719
13.9988 | -105.856
14.0025 | -105.839
14.0062 | -105.8192
14.0098 | -105.7953
14.0135 | -105.7688
14.0172 | -105.7448
14.0208 | -105.7227
14.0245 | -105.6991
14.0282 | -105.6748
14.0319 | -105.6483
14.0355 | -105.6208
14.0392 | -105.597
14.0429 | -105.5754
14.0466 | -105.5549
14.0502 | -105.5346
14.0539 | -105.5181
14.0576 | -105.4994
14.0613 | -105.4827
14.0649 | -105.472
14.0686 | -105.4646
14.0723 | -105.455
14.0759 | -105.4477
14.0796 | -105.4428
14.0833 | -105.442
14.087 | -105.4454
14.0906 | -105.4492
14.0943 | -105.4551
14.098 | -105.4625
14.1017 | -105.469
14.1053 | -105.477
14.109 | -105.4859
14.1127 | -105.4959
14.1164 | -105.5096
14.12 | -105.524
14.1237 | -105.5395
14.1274 | -105.5558
14.131 | -105.5733
14.1347 | -105.59
14.1384 | -105.6099
14.1421 | -105.6258
14.1457 | -105.6419
14.1494 | -105.6592
14.1531 | -105.6764
14.1568 | -105.6962
14.1604 | -105.7256
14.1641 | -105.7588
14.1678 | -105.7881
14.1715 | -105.8113
14.1751 | -105.834
14.1788 | -105.8601
14.1825 | -105.8883
14.1861 | -105.9164
14.1898 | -105.9421
14.1935 | -105.9703
14.1972 | -105.9953
14.2008 | -106.0186
14.2045 | -106.044
14.2082 | -106.0724
14.2119 | -106.1006
14.2155 | -106.1265
14.2192 | -106.1478
14.2229 | -106.1657
14.2265 | -106.1825
14.2302 | -106.199
14.2339 | -106.2151
14.2376 | -106.2281
14.2412 | -106.2365
14.2449 | -106.245
14.2486 | -106.2541
14.2523 | -106.2617
14.2559 | -106.2629
14.2596 | -106.2662
14.2633 | -106.2645
14.267 | -106.259
14.2706 | -106.2516
14.2743 | -106.2413
14.278 | -106.2301
14.2816 | -106.2161
14.2853 | -106.197
14.289 | -106.1753
14.2927 | -106.1562
14.2963 | -106.1343
14.3 | -106.1109
14.3037 | -106.0831
14.3074 | -106.0514
14.311 | -106.0262
14.3147 | -106.0004
14.3184 | -105.9765
14.3221 | -105.9508
14.3257 | -105.9263
14.3294 | -105.9015
14.3331 | -105.8784
14.3367 | -105.8581
14.3404 | -105.8397
14.3441 | -105.8227
14.3478 | -105.8119
14.3514 | -105.802
14.3551 | -105.7917
14.3588 | -105.7844
14.3625 | -105.7798
14.3661 | -105.7739
14.3698 | -105.7712
14.3735 | -105.7728
14.3772 | -105.7787
14.3808 | -105.7848
14.3845 | -105.7898
14.3882 | -105.7966
14.3918 | -105.8078
14.3955 | -105.8223
14.3992 | -105.8347
14.4029 | -105.8501
14.4065 | -105.8642
14.4102 | -105.8773
14.4139 | -105.8888
14.4176 | -105.9039
14.4212 | -105.9179
14.4249 | -105.9315
14.4286 | -105.9474
14.4323 | -105.9623
14.4359 | -105.9747
14.4396 | -105.9856
14.4433 | -106.0012
14.4469 | -106.0149
14.4506 | -106.0251
14.4543 | -106.0349
14.458 | -106.0435
14.4616 | -106.0556
14.4653 | -106.0667
14.469 | -106.0782
14.4727 | -106.0889
14.4763 | -106.0987
14.48 | -106.1094
14.4837 | -106.1146
14.4873 | -106.1195
14.491 | -106.1261
14.4947 | -106.1312
14.4984 | -106.1331
14.502 | -106.1352
14.5057 | -106.1381
14.5094 | -106.1426
14.5131 | -106.1466
14.5167 | -106.149
14.5204 | -106.152
14.5241 | -106.1536
14.5278 | -106.1508
14.5314 | -106.1517
14.5351 | -106.1536
14.5388 | -106.1551
14.5424 | -106.1571
14.5461 | -106.1585
14.5498 | -106.1594
14.5535 | -106.1616
14.5571 | -106.1631
14.5608 | -106.1637
14.5645 | -106.1679
14.5682 | -106.1753
14.5718 | -106.1812
14.5755 | -106.1841
14.5792 | -106.1863
14.5829 | -106.1912
14.5865 | -106.1987
14.5902 | -106.2076
14.5939 | -106.2168
14.5975 | -106.2272
14.6012 | -106.2384
14.6049 | -106.2475
14.6086 | -106.2595
14.6122 | -106.2722
14.6159 | -106.2879
14.6196 | -106.3048
14.6233 | -106.3211
14.6269 | -106.3347
14.6306 | -106.3534
14.6343 | -106.3718
14.638 | -106.3873
14.6416 | -106.4076
14.6453 | -106.4268
14.649 | -106.4454
14.6526 | -106.463
14.6563 | -106.4811
14.66 | -106.5024
14.6637 | -106.5185
14.6673 | -106.5369
14.671 | -106.553
14.6747 | -106.5654
14.6784 | -106.5771
14.682 | -106.5866
14.6857 | -106.5944
14.6894 | -106.6015
14.6931 | -106.6115
14.6967 | -106.6171
14.7004 | -106.6205
14.7041 | -106.6224
14.7077 | -106.6236
14.7114 | -106.6246
14.7151 | -106.6225
14.7188 | -106.6198
14.7224 | -106.6144
14.7261 | -106.6036
14.7298 | -106.5927
14.7335 | -106.5877
14.7371 | -106.5844
14.7408 | -106.5732
14.7445 | -106.5633
14.7482 | -106.5547
14.7518 | -106.5454
14.7555 | -106.5335
14.7592 | -106.5194
14.7628 | -106.5042
14.7665 | -106.4903
14.7702 | -106.4779
14.7739 | -106.4638
14.7775 | -106.4486
14.7812 | -106.4331
14.7849 | -106.4204
14.7886 | -106.4102
14.7922 | -106.3958
14.7959 | -106.3821
14.7996 | -106.3698
14.8033 | -106.3593
14.8069 | -106.3486
14.8106 | -106.3372
14.8143 | -106.325
14.8179 | -106.311
14.8216 | -106.2984
14.8253 | -106.287
14.829 | -106.2753
14.8326 | -106.2656
14.8363 | -106.2599
14.84 | -106.2547
14.8437 | -106.248
14.8473 | -106.24
14.851 | -106.2339
14.8547 | -106.2255
14.8583 | -106.2172
14.862 | -106.2114
14.8657 | -106.2068
14.8694 | -106.2048
14.873 | -106.1994
14.8767 | -106.1929
14.8804 | -106.188
14.8841 | -106.1864
14.8877 | -106.185
14.8914 | -106.1831
14.8951 | -106.1827
14.8988 | -106.1811
14.9024 | -106.1756
14.9061 | -106.17
14.9098 | -106.1668
14.9134 | -106.1625
14.9171 | -106.1579
14.9208 | -106.1542
14.9245 | -106.1489
14.9281 | -106.1415
14.9318 | -106.1348
14.9355 | -106.128
14.9392 | -106.1225
14.9428 | -106.115
14.9465 | -106.1035
14.9502 | -106.0941
14.9539 | -106.0892
14.9575 | -106.0801
14.9612 | -106.0712
14.9649 | -106.0628
14.9685 | -106.0552
14.9722 | -106.0505
14.9759 | -106.0454
14.9796 | -106.0417
14.9832 | -106.0391
14.9869 | -106.0379
14.9906 | -106.037
14.9943 | -106.0387
14.9979 | -106.0424
15.0016 | -106.0462
15.0053 | -106.0556
15.009 | -106.0657
15.0126 | -106.0792
15.0163 | -106.0962
15.02 | -106.1158
15.0236 | -106.1377
15.0273 | -106.1634
15.031 | -106.1915
15.0347 | -106.2199
15.0383 | -106.2483
15.042 | -106.2787
15.0457 | -106.3106
15.0494 | -106.3429
15.053 | -106.3746
15.0567 | -106.4055
15.0604 | -106.4385
15.0641 | -106.4738
15.0677 | -106.5056
15.0714 | -106.5397
15.0751 | -106.5733
15.0787 | -106.6013
15.0824 | -106.6219
15.0861 | -106.6388
15.0898 | -106.6607
15.0934 | -106.684
15.0971 | -106.701
15.1008 | -106.7123
15.1045 | -106.7197
15.1081 | -106.7272
15.1118 | -106.7302
15.1155 | -106.7308
15.1192 | -106.7317
15.1228 | -106.7287
15.1265 | -106.7232
15.1302 | -106.7166
15.1338 | -106.7072
15.1375 | -106.6908
15.1412 | -106.6719
15.1449 | -106.6569
15.1485 | -106.6405
15.1522 | -106.6268
15.1559 | -106.6107
15.1596 | -106.5937
15.1632 | -106.577
15.1669 | -106.5588
15.1706 | -106.5389
15.1743 | -106.5217
15.1779 | -106.5081
15.1816 | -106.497
15.1853 | -106.4907
15.1889 | -106.4825
15.1926 | -106.4753
15.1963 | -106.4686
15.2 | -106.4636
15.2036 | -106.4602
15.2073 | -106.4576
15.211 | -106.4572
15.2147 | -106.4597
15.2183 | -106.4629
15.222 | -106.4628
15.2257 | -106.4612
15.2293 | -106.4597
15.233 | -106.4588
15.2367 | -106.4571
15.2404 | -106.4562
15.244 | -106.4547
15.2477 | -106.4516
15.2514 | -106.4475
15.2551 | -106.4412
15.2587 | -106.4313
15.2624 | -106.4205
15.2661 | -106.4125
15.2698 | -106.4051
15.2734 | -106.3963
15.2771 | -106.3855
15.2808 | -106.3744
15.2844 | -106.3639
15.2881 | -106.3511
15.2918 | -106.3391
15.2955 | -106.327
15.2991 | -106.3195
15.3028 | -106.3139
15.3065 | -106.3037
15.3102 | -106.2941
15.3138 | -106.288
15.3175 | -106.2828
15.3212 | -106.2804
15.3249 | -106.2774
15.3285 | -106.2763
15.3322 | -106.2779
15.3359 | -106.2813
15.3395 | -106.2864
15.3432 | -106.2922
15.3469 | -106.2988
15.3506 | -106.308
15.3542 | -106.3225
15.3579 | -106.34
15.3616 | -106.3548
15.3653 | -106.3729
15.3689 | -106.3874
15.3726 | -106.4017
15.3763 | -106.4203
15.38 | -106.4388
15.3836 | -106.4576
15.3873 | -106.4774
15.391 | -106.4973
15.3946 | -106.5147
15.3983 | -106.53
15.402 | -106.5453
15.4057 | -106.5563
15.4093 | -106.5693
15.413 | -106.583
15.4167 | -106.594
15.4204 | -106.6038
15.424 | -106.6127
15.4277 | -106.6199
15.4314 | -106.6239
15.4351 | -106.6246
15.4387 | -106.6236
15.4424 | -106.6201
15.4461 | -106.6167
15.4497 | -106.6111
15.4534 | -106.6047
15.4571 | -106.5965
15.4608 | -106.59
15.4644 | -106.5802
15.4681 | -106.568
15.4718 | -106.553
15.4755 | -106.5409
15.4791 | -106.5293
15.4828 | -106.5158
15.4865 | -106.5008
15.4902 | -106.4839
15.4938 | -106.4678
15.4975 | -106.4516
15.5012 | -106.4362
15.5048 | -106.4221
15.5085 | -106.4074
15.5122 | -106.3921
15.5159 | -106.3782
15.5195 | -106.3641
15.5232 | -106.3502
15.5269 | -106.3354
15.5306 | -106.3211
15.5342 | -106.3065
15.5379 | -106.2943
15.5416 | -106.2846
15.5452 | -106.275
15.5489 | -106.267
15.5526 | -106.2603
15.5563 | -106.2532
15.5599 | -106.2458
15.5636 | -106.2402
15.5673 | -106.2349
15.571 | -106.2328
15.5746 | -106.2318
15.5783 | -106.2276
15.582 | -106.2265
15.5857 | -106.2282
15.5893 | -106.2326
15.593 | -106.2374
15.5967 | -106.2428
15.6003 | -106.2472
15.604 | -106.2524
15.6077 | -106.2613
15.6114 | -106.2699
15.615 | -106.2811
15.6187 | -106.295
15.6224 | -106.3074
15.6261 | -106.3195
15.6297 | -106.3341
15.6334 | -106.3492
15.6371 | -106.3647
15.6408 | -106.3811
15.6444 | -106.3944
15.6481 | -106.4095
15.6518 | -106.4275
15.6554 | -106.4449
15.6591 | -106.4603
15.6628 | -106.4764
15.6665 | -106.4951
15.6701 | -106.5145
15.6738 | -106.5321
15.6775 | -106.5469
15.6812 | -106.5636
15.6848 | -106.5754
15.6885 | -106.5905
15.6922 | -106.6063
15.6959 | -106.6203
15.6995 | -106.6331
15.7032 | -106.645
15.7069 | -106.6524
15.7105 | -106.6591
15.7142 | -106.6671
15.7179 | -106.6736
15.7216 | -106.6765
15.7252 | -106.6781
15.7289 | -106.6786
15.7326 | -106.6765
15.7363 | -106.6736
15.7399 | -106.6695
15.7436 | -106.6619
15.7473 | -106.6538
15.751 | -106.6464
15.7546 | -106.6374
15.7583 | -106.6285
15.762 | -106.6159
15.7656 | -106.6035
15.7693 | -106.5926
15.773 | -106.5804
15.7767 | -106.567
15.7803 | -106.5556
15.784 | -106.5467
15.7877 | -106.5346
15.7914 | -106.5249
15.795 | -106.5129
15.7987 | -106.5022
15.8024 | -106.4978
15.8061 | -106.4886
15.8097 | -106.4794
15.8134 | -106.4704
15.8171 | -106.4624
15.8207 | -106.4575
15.8244 | -106.4543
15.8281 | -106.453
15.8318 | -106.4499
15.8354 | -106.4478
15.8391 | -106.4457
15.8428 | -106.4432
15.8465 | -106.4421
15.8501 | -106.441
15.8538 | -106.4399
15.8575 | -106.439
15.8612 | -106.437
15.8648 | -106.4357
15.8685 | -106.4327
15.8722 | -106.4262
15.8758 | -106.422
15.8795 | -106.4204
15.8832 | -106.4173
15.8869 | -106.4116
15.8905 | -106.4041
15.8942 | -106.3969
15.8979 | -106.3902
15.9016 | -106.3845
15.9052 | -106.3813
15.9089 | -106.3774
15.9126 | -106.3742
15.9162 | -106.3681
15.9199 | -106.3611
15.9236 | -106.3562
15.9273 | -106.3517
15.9309 | -106.3469
15.9346 | -106.3462
15.9383 | -106.3439
15.942 | -106.3408
15.9456 | -106.3401
15.9493 | -106.3418
15.953 | -106.342
15.9567 | -106.3427
15.9603 | -106.3468
15.964 | -106.348
15.9677 | -106.3496
15.9713 | -106.3543
15.975 | -106.3582
15.9787 | -106.3639
15.9824 | -106.3682
15.986 | -106.3733
15.9897 | -106.3813
15.9934 | -106.3892
15.9971 | -106.3975
16.0007 | -106.4055
16.0044 | -106.4128
16.0081 | -106.4191
16.0118 | -106.4235
16.0154 | -106.4305
16.0191 | -106.4364
16.0228 | -106.447
16.0264 | -106.4537
16.0301 | -106.457
16.0338 | -106.4633
16.0375 | -106.4686
16.0411 | -106.475
16.0448 | -106.4795
16.0485 | -106.4822
16.0522 | -106.4851
16.0558 | -106.4897
16.0595 | -106.4938
16.0632 | -106.4942
16.0669 | -106.4966
16.0705 | -106.5001
16.0742 | -106.5029
16.0779 | -106.5028
16.0815 | -106.4993
16.0852 | -106.4962
16.0889 | -106.4943
16.0926 | -106.492
16.0962 | -106.4922
16.0999 | -106.4919
16.1036 | -106.4879
16.1073 | -106.4823
16.1109 | -106.4793
16.1146 | -106.4767
16.1183 | -106.4728
16.122 | -106.4666
16.1256 | -106.4617
16.1293 | -106.4585
16.133 | -106.4539
16.1366 | -106.4477
16.1403 | -106.4416
16.144 | -106.439
16.1477 | -106.4387
16.1513 | -106.4376
16.155 | -106.4354
16.1587 | -106.4305
16.1624 | -106.4289
16.166 | -106.4262
16.1697 | -106.424
16.1734 | -106.4217
16.1771 | -106.4177
16.1807 | -106.4139
16.1844 | -106.4117
16.1881 | -106.4103
16.1917 | -106.4088
16.1954 | -106.4111
16.1991 | -106.4126
16.2028 | -106.4122
16.2064 | -106.413
16.2101 | -106.4145
16.2138 | -106.4192
16.2175 | -106.4234
16.2211 | -106.4278
16.2248 | -106.4317
16.2285 | -106.4378
16.2322 | -106.4467
16.2358 | -106.4568
16.2395 | -106.4653
16.2432 | -106.4765
16.2468 | -106.4868
16.2505 | -106.4956
16.2542 | -106.5026
16.2579 | -106.5125
16.2615 | -106.5238
16.2652 | -106.5343
16.2689 | -106.5429
16.2726 | -106.5492
16.2762 | -106.555
16.2799 | -106.561
16.2836 | -106.5672
16.2872 | -106.5694
16.2909 | -106.5723
16.2946 | -106.5803
16.2983 | -106.5841
16.3019 | -106.5881
16.3056 | -106.5922
16.3093 | -106.5941
16.313 | -106.5974
16.3166 | -106.5996
16.3203 | -106.5998
16.324 | -106.5999
16.3277 | -106.6007
16.3313 | -106.6022
16.335 | -106.6005
16.3387 | -106.602
16.3423 | -106.6029
16.346 | -106.6031
16.3497 | -106.6041
16.3534 | -106.6048
16.357 | -106.6055
16.3607 | -106.6063
16.3644 | -106.6074
16.3681 | -106.606
16.3717 | -106.6058
16.3754 | -106.6073
16.3791 | -106.6069
16.3828 | -106.6056
16.3864 | -106.6049
16.3901 | -106.6064
16.3938 | -106.6061
16.3974 | -106.6047
16.4011 | -106.6015
16.4048 | -106.5981
16.4085 | -106.5976
16.4121 | -106.5988
16.4158 | -106.5976
16.4195 | -106.5959
16.4232 | -106.5926
16.4268 | -106.5897
16.4305 | -106.5864
16.4342 | -106.5854
16.4379 | -106.5848
16.4415 | -106.5861
16.4452 | -106.5861
16.4489 | -106.5847
16.4525 | -106.5849
16.4562 | -106.584
16.4599 | -106.5845
16.4636 | -106.5838
16.4672 | -106.583
16.4709 | -106.585
16.4746 | -106.5895
16.4783 | -106.5926
16.4819 | -106.5959
16.4856 | -106.5979
16.4893 | -106.5994
16.493 | -106.6022
16.4966 | -106.6044
16.5003 | -106.608
16.504 | -106.6126
16.5076 | -106.6166
16.5113 | -106.6194
16.515 | -106.6225
16.5187 | -106.6255
16.5223 | -106.6286
16.526 | -106.6308
16.5297 | -106.6329
16.5334 | -106.638
16.537 | -106.6422
16.5407 | -106.6469
16.5444 | -106.6547
16.5481 | -106.6619
16.5517 | -106.6653
16.5554 | -106.6693
16.5591 | -106.6722
16.5627 | -106.6725
16.5664 | -106.675
16.5701 | -106.6816
16.5738 | -106.6858
16.5774 | -106.6891
16.5811 | -106.6956
16.5848 | -106.7024
16.5885 | -106.7091
16.5921 | -106.7148
16.5958 | -106.7191
16.5995 | -106.723
16.6031 | -106.7321
16.6068 | -106.7424
16.6105 | -106.7498
16.6142 | -106.7592
16.6178 | -106.7689
16.6215 | -106.7792
16.6252 | -106.7878
16.6289 | -106.7952
16.6325 | -106.804
16.6362 | -106.8156
16.6399 | -106.8279
16.6436 | -106.8384
16.6472 | -106.848
16.6509 | -106.8603
16.6546 | -106.8718
16.6582 | -106.8848
16.6619 | -106.8981
16.6656 | -106.9103
16.6693 | -106.9235
16.6729 | -106.9341
16.6766 | -106.9431
16.6803 | -106.9489
16.684 | -106.9556
16.6876 | -106.9627
16.6913 | -106.9697
16.695 | -106.9795
16.6987 | -106.9862
16.7023 | -106.9897
16.706 | -106.992
16.7097 | -106.9916
16.7133 | -106.9896
16.717 | -106.9873
16.7207 | -106.9859
16.7244 | -106.9858
16.728 | -106.9825
16.7317 | -106.9774
16.7354 | -106.9728
16.7391 | -106.9683
16.7427 | -106.9621
16.7464 | -106.9556
16.7501 | -106.9518
16.7538 | -106.9458
16.7574 | -106.9383
16.7611 | -106.9323
16.7648 | -106.9249
16.7684 | -106.9168
16.7721 | -106.9071
16.7758 | -106.8962
16.7795 | -106.8834
16.7831 | -106.8728
16.7868 | -106.8617
16.7905 | -106.8552
16.7942 | -106.8492
16.7978 | -106.8434
16.8015 | -106.8377
16.8052 | -106.8305
16.8089 | -106.823
16.8125 | -106.8167
16.8162 | -106.8131
16.8199 | -106.8083
16.8235 | -106.8054
16.8272 | -106.804
16.8309 | -106.8042
16.8346 | -106.806
16.8382 | -106.8107
16.8419 | -106.8148
16.8456 | -106.817
16.8493 | -106.8199
16.8529 | -106.8271
16.8566 | -106.8365
16.8603 | -106.8458
16.864 | -106.857
16.8676 | -106.8707
16.8713 | -106.8859
16.875 | -106.8997
16.8786 | -106.9143
16.8823 | -106.928
16.886 | -106.941
16.8897 | -106.9541
16.8933 | -106.9681
16.897 | -106.9823
16.9007 | -106.9965
16.9044 | -107.0119
16.908 | -107.0256
16.9117 | -107.0393
16.9154 | -107.0541
16.919 | -107.0684
16.9227 | -107.0797
16.9264 | -107.091
16.9301 | -107.0972
16.9337 | -107.1022
16.9374 | -107.1107
16.9411 | -107.1209
16.9448 | -107.1294
16.9484 | -107.1349
16.9521 | -107.1392
16.9558 | -107.1393
16.9595 | -107.1436
16.9631 | -107.1464
16.9668 | -107.1444
16.9705 | -107.1424
16.9741 | -107.1405
16.9778 | -107.1375
16.9815 | -107.1364
16.9852 | -107.135
16.9888 | -107.1329
16.9925 | -107.1283
16.9962 | -107.1214
16.9999 | -107.1151
17.0035 | -107.1072
17.0072 | -107.0998
17.0109 | -107.0925
17.0146 | -107.0824
17.0182 | -107.073
17.0219 | -107.0659
17.0256 | -107.0561
17.0292 | -107.0462
17.0329 | -107.0356
17.0366 | -107.0246
17.0403 | -107.0127
17.0439 | -107.0001
17.0476 | -106.9835
17.0513 | -106.9688
17.055 | -106.9541
17.0586 | -106.9387
17.0623 | -106.9216
17.066 | -106.9048
17.0697 | -106.8848
17.0733 | -106.865
17.077 | -106.8484
17.0807 | -106.8286
17.0843 | -106.8065
17.088 | -106.7871
17.0917 | -106.7685
17.0954 | -106.7499
17.099 | -106.7311
17.1027 | -106.7109
17.1064 | -106.6898
17.1101 | -106.6696
17.1137 | -106.6499
17.1174 | -106.6326
17.1211 | -106.6153
17.1248 | -106.5995
17.1284 | -106.5866
17.1321 | -106.5742
17.1358 | -106.5618
17.1394 | -106.5498
17.1431 | -106.5398
17.1468 | -106.5312
17.1505 | -106.5233
17.1541 | -106.5184
17.1578 | -106.5147
17.1615 | -106.514
17.1652 | -106.515
17.1688 | -106.5145
17.1725 | -106.5175
17.1762 | -106.5209
17.1799 | -106.5237
17.1835 | -106.5299
17.1872 | -106.5388
17.1909 | -106.5488
17.1945 | -106.5604
17.1982 | -106.5709
17.2019 | -106.5829
17.2056 | -106.5978
17.2092 | -106.6147
17.2129 | -106.6363
17.2166 | -106.6576
17.2203 | -106.6778
17.2239 | -106.6967
17.2276 | -106.7133
17.2313 | -106.7319
17.235 | -106.7507
17.2386 | -106.7719
17.2423 | -106.7943
17.246 | -106.8142
17.2496 | -106.8329
17.2533 | -106.85
17.257 | -106.8667
17.2607 | -106.8841
17.2643 | -106.9001
17.268 | -106.9116
17.2717 | -106.9209
17.2754 | -106.9302
17.279 | -106.9374
17.2827 | -106.9396
17.2864 | -106.9418
17.2901 | -106.9412
17.2937 | -106.9409
17.2974 | -106.9377
17.3011 | -106.9298
17.3047 | -106.9195
17.3084 | -106.9108
17.3121 | -106.8949
17.3158 | -106.8766
17.3194 | -106.8588
17.3231 | -106.8363
17.3268 | -106.8144
17.3305 | -106.7909
17.3341 | -106.7648
17.3378 | -106.7366
17.3415 | -106.7038
17.3451 | -106.6712
17.3488 | -106.6402
17.3525 | -106.6112
17.3562 | -106.58
17.3598 | -106.5488
17.3635 | -106.5173
17.3672 | -106.4873
17.3709 | -106.4596
17.3745 | -106.4327
17.3782 | -106.4082
17.3819 | -106.3859
17.3856 | -106.3651
17.3892 | -106.3456
17.3929 | -106.3293
17.3966 | -106.3169
17.4002 | -106.3065
17.4039 | -106.2986
17.4076 | -106.2925
17.4113 | -106.2911
17.4149 | -106.2938
17.4186 | -106.2973
17.4223 | -106.3049
17.426 | -106.3163
17.4296 | -106.3299
17.4333 | -106.3451
17.437 | -106.3638
17.4407 | -106.3841
17.4443 | -106.4072
17.448 | -106.4311
17.4517 | -106.4564
17.4553 | -106.4821
17.459 | -106.5094
17.4627 | -106.5379
17.4664 | -106.5664
17.47 | -106.5964
17.4737 | -106.6276
17.4774 | -106.6568
17.4811 | -106.6875
17.4847 | -106.7093
17.4884 | -106.7351
17.4921 | -106.7614
17.4958 | -106.7846
17.4994 | -106.8071
17.5031 | -106.8291
17.5068 | -106.846
17.5104 | -106.8641
17.5141 | -106.8806
17.5178 | -106.8964
17.5215 | -106.9111
17.5251 | -106.9206
17.5288 | -106.9267
17.5325 | -106.9332
17.5362 | -106.9374
17.5398 | -106.9417
17.5435 | -106.9441
17.5472 | -106.9439
17.5509 | -106.9423
17.5545 | -106.937
17.5582 | -106.9303
17.5619 | -106.9262
17.5655 | -106.9214
17.5692 | -106.9125
17.5729 | -106.9036
17.5766 | -106.8957
17.5802 | -106.8886
17.5839 | -106.8795
17.5876 | -106.8718
17.5913 | -106.866
17.5949 | -106.8594
17.5986 | -106.8525
17.6023 | -106.8433
17.606 | -106.8352
17.6096 | -106.8273
17.6133 | -106.8213
17.617 | -106.8174
17.6206 | -106.813
17.6243 | -106.8072
17.628 | -106.8024
17.6317 | -106.7981
17.6353 | -106.7948
17.639 | -106.7924
17.6427 | -106.7908
17.6464 | -106.7905
17.65 | -106.7898
17.6537 | -106.7884
17.6574 | -106.786
17.661 | -106.784
17.6647 | -106.785
17.6684 | -106.7891
17.6721 | -106.7918
17.6757 | -106.7932
17.6794 | -106.7942
17.6831 | -106.7954
17.6868 | -106.7957
17.6904 | -106.7983
17.6941 | -106.8016
17.6978 | -106.8051
17.7015 | -106.805
17.7051 | -106.8074
17.7088 | -106.8089
17.7125 | -106.8113
17.7161 | -106.8136
17.7198 | -106.8149
17.7235 | -106.8141
17.7272 | -106.8138
17.7308 | -106.8118
17.7345 | -106.8093
17.7382 | -106.8088
17.7419 | -106.8061
17.7455 | -106.8032
17.7492 | -106.8041
17.7529 | -106.8041
17.7566 | -106.8026
17.7602 | -106.8038
17.7639 | -106.8015
17.7676 | -106.7969
17.7712 | -106.7953
17.7749 | -106.7925
17.7786 | -106.7918
17.7823 | -106.7922
17.7859 | -106.7931
17.7896 | -106.7948
17.7933 | -106.7953
17.797 | -106.7955
17.8006 | -106.7939
17.8043 | -106.7966
17.808 | -106.8017
17.8117 | -106.8048
17.8153 | -106.8082
17.819 | -106.8123
17.8227 | -106.815
17.8263 | -106.8178
17.83 | -106.8179
17.8337 | -106.821
17.8374 | -106.8236
17.841 | -106.8247
17.8447 | -106.8255
17.8484 | -106.8286
17.8521 | -106.8287
17.8557 | -106.8305
17.8594 | -106.8299
17.8631 | -106.8302
17.8668 | -106.8299
17.8704 | -106.8281
17.8741 | -106.8267
17.8778 | -106.8256
17.8814 | -106.823
17.8851 | -106.8229
17.8888 | -106.8228
17.8925 | -106.8204
17.8961 | -106.8162
17.8998 | -106.8105
17.9035 | -106.8063
17.9072 | -106.8038
17.9108 | -106.8011
17.9145 | -106.7957
17.9182 | -106.7868
17.9219 | -106.7786
17.9255 | -106.7722
17.9292 | -106.7668
17.9329 | -106.7616
17.9365 | -106.7583
17.9402 | -106.7557
17.9439 | -106.7509
17.9476 | -106.7466
17.9512 | -106.7424
17.9549 | -106.7397
17.9586 | -106.7375
17.9623 | -106.7375
17.9659 | -106.7357
17.9696 | -106.7364
17.9733 | -106.7372
17.9769 | -106.7357
17.9806 | -106.7335
17.9843 | -106.7346
17.988 | -106.7368
17.9916 | -106.7384
17.9953 | -106.7415
17.999 | -106.7426
18.0027 | -106.7405
18.0063 | -106.7403
18.01 | -106.7409
18.0137 | -106.7411
18.0174 | -106.7381
18.021 | -106.7366
18.0247 | -106.7373
18.0284 | -106.7336
18.032 | -106.7305
18.0357 | -106.7293
18.0394 | -106.7277
18.0431 | -106.7239
18.0467 | -106.7174
18.0504 | -106.7122
18.0541 | -106.7081
18.0578 | -106.705
18.0614 | -106.7
18.0651 | -106.6955
18.0688 | -106.688
18.0725 | -106.6772
18.0761 | -106.6681
18.0798 | -106.6593
18.0835 | -106.6487
18.0871 | -106.6403
18.0908 | -106.6287
18.0945 | -106.6168
18.0982 | -106.6054
18.1018 | -106.5931
18.1055 | -106.5805
18.1092 | -106.5707
18.1129 | -106.5592
18.1165 | -106.5518
18.1202 | -106.5419
18.1239 | -106.5311
18.1276 | -106.5216
18.1312 | -106.5122
18.1349 | -106.5048
18.1386 | -106.4971
18.1422 | -106.4886
18.1459 | -106.4819
18.1496 | -106.4779
18.1533 | -106.4743
18.1569 | -106.4709
18.1606 | -106.4688
18.1643 | -106.4688
18.168 | -106.4686
18.1716 | -106.471
18.1753 | -106.4745
18.179 | -106.4765
18.1827 | -106.4806
18.1863 | -106.4868
18.19 | -106.4918
18.1937 | -106.4978
18.1973 | -106.5049
18.201 | -106.5143
18.2047 | -106.5249
18.2084 | -106.5371
18.212 | -106.5498
18.2157 | -106.5616
18.2194 | -106.5725
18.2231 | -106.5847
18.2267 | -106.5969
18.2304 | -106.6103
18.2341 | -106.6241
18.2378 | -106.6397
18.2414 | -106.6537
18.2451 | -106.6658
18.2488 | -106.6781
18.2524 | -106.6884
18.2561 | -106.6993
18.2598 | -106.7133
18.2635 | -106.7221
18.2671 | -106.7304
18.2708 | -106.7373
18.2745 | -106.7451
18.2782 | -106.7519
18.2818 | -106.7537
18.2855 | -106.7574
18.2892 | -106.7621
18.2929 | -106.7648
18.2965 | -106.7651
18.3002 | -106.7645
18.3039 | -106.7636
18.3075 | -106.7622
18.3112 | -106.7584
18.3149 | -106.7502
18.3186 | -106.7412
18.3222 | -106.733
18.3259 | -106.7264
18.3296 | -106.7195
18.3333 | -106.7151
18.3369 | -106.7061
18.3406 | -106.697
18.3443 | -106.6895
18.348 | -106.6817
18.3516 | -106.6736
18.3553 | -106.6675
18.359 | -106.6609
18.3626 | -106.6558
18.3663 | -106.6482
18.37 | -106.6406
18.3737 | -106.6326
18.3773 | -106.6244
18.381 | -106.6176
18.3847 | -106.6129
18.3884 | -106.6078
18.392 | -106.6036
18.3957 | -106.5992
18.3994 | -106.5972
18.403 | -106.5958
18.4067 | -106.5941
18.4104 | -106.5926
18.4141 | -106.59
18.4177 | -106.5897
18.4214 | -106.5897
18.4251 | -106.5885
18.4288 | -106.5905
18.4324 | -106.593
18.4361 | -106.5959
18.4398 | -106.5995
18.4435 | -106.6002
18.4471 | -106.6029
18.4508 | -106.6096
18.4545 | -106.6162
18.4581 | -106.622
18.4618 | -106.6264
18.4655 | -106.6335
18.4692 | -106.6394
18.4728 | -106.6454
18.4765 | -106.6536
18.4802 | -106.6631
18.4839 | -106.6746
18.4875 | -106.6856
18.4912 | -106.6946
18.4949 | -106.7058
18.4986 | -106.7139
18.5022 | -106.7213
18.5059 | -106.7299
18.5096 | -106.7379
18.5132 | -106.7465
18.5169 | -106.7537
18.5206 | -106.7574
18.5243 | -106.7619
18.5279 | -106.7646
18.5316 | -106.7669
18.5353 | -106.7679
18.539 | -106.7677
18.5426 | -106.7696
18.5463 | -106.7624
18.55 | -106.758
18.5537 | -106.7521
18.5573 | -106.7463
18.561 | -106.7419
18.5647 | -106.7364
18.5683 | -106.7302
18.572 | -106.7236
18.5757 | -106.7175
18.5794 | -106.7124
18.583 | -106.7072
18.5867 | -106.6985
18.5904 | -106.6866
18.5941 | -106.6771
18.5977 | -106.6689
18.6014 | -106.6629
18.6051 | -106.6546
18.6088 | -106.6462
18.6124 | -106.6382
18.6161 | -106.6312
18.6198 | -106.6233
18.6234 | -106.6163
18.6271 | -106.6074
18.6308 | -106.598
18.6345 | -106.5906
18.6381 | -106.5861
18.6418 | -106.5821
18.6455 | -106.5781
18.6492 | -106.5739
18.6528 | -106.5704
18.6565 | -106.5675
18.6602 | -106.5652
18.6639 | -106.5636
18.6675 | -106.5616
18.6712 | -106.56
18.6749 | -106.5595
18.6785 | -106.5629
18.6822 | -106.5654
18.6859 | -106.5675
18.6896 | -106.5703
18.6932 | -106.572
18.6969 | -106.5733
18.7006 | -106.5777
18.7043 | -106.5846
18.7079 | -106.5912
18.7116 | -106.5956
18.7153 | -106.6004
18.7189 | -106.6063
18.7226 | -106.6144
18.7263 | -106.6219
18.73 | -106.6263
18.7336 | -106.6336
18.7373 | -106.6429
18.741 | -106.651
18.7447 | -106.6592
18.7483 | -106.6677
18.752 | -106.676
18.7557 | -106.6847
18.7594 | -106.6931
18.763 | -106.7039
18.7667 | -106.7117
18.7704 | -106.7218
18.774 | -106.7331
18.7777 | -106.7427
18.7814 | -106.7548
18.7851 | -106.7677
18.7887 | -106.7777
18.7924 | -106.7871
18.7961 | -106.799
18.7998 | -106.8097
18.8034 | -106.8199
18.8071 | -106.8301
18.8108 | -106.8387
18.8145 | -106.8478
18.8181 | -106.8584
18.8218 | -106.8677
18.8255 | -106.8767
18.8291 | -106.8826
18.8328 | -106.8895
18.8365 | -106.8983
18.8402 | -106.9081
18.8438 | -106.9167
18.8475 | -106.9252
18.8512 | -106.9325
18.8549 | -106.942
18.8585 | -106.9498
18.8622 | -106.9578
18.8659 | -106.9627
18.8696 | -106.9672
18.8732 | -106.974
18.8769 | -106.9768
18.8806 | -106.9784
18.8842 | -106.9802
18.8879 | -106.9829
18.8916 | -106.9834
18.8953 | -106.9862
18.8989 | -106.9892
18.9026 | -106.9902
18.9063 | -106.9895
18.91 | -106.9886
18.9136 | -106.9897
18.9173 | -106.9908
18.921 | -106.9903
18.9247 | -106.9871
18.9283 | -106.9867
18.932 | -106.9864
18.9357 | -106.9849
18.9393 | -106.9841
18.943 | -106.9807
18.9467 | -106.9773
18.9504 | -106.9723
18.954 | -106.9678
18.9577 | -106.9645
18.9614 | -106.961
18.9651 | -106.9563
18.9687 | -106.9499
18.9724 | -106.9418
18.9761 | -106.9343
18.9798 | -106.9286
18.9834 | -106.9208
18.9871 | -106.9133
18.9908 | -106.9048
18.9944 | -106.8971
18.9981 | -106.89
19.0018 | -106.8808
19.0055 | -106.8707
19.0091 | -106.8596
19.0128 | -106.8481
19.0165 | -106.8365
19.0202 | -106.8256
19.0238 | -106.8158
19.0275 | -106.8065
19.0312 | -106.7976
19.0348 | -106.7906
19.0385 | -106.7829
19.0422 | -106.7759
19.0459 | -106.7704
19.0495 | -106.7636
19.0532 | -106.7568
19.0569 | -106.7533
19.0606 | -106.7497
19.0642 | -106.7438
19.0679 | -106.7402
19.0716 | -106.7391
19.0753 | -106.7386
19.0789 | -106.7368
19.0826 | -106.7369
19.0863 | -106.7365
19.0899 | -106.7348
19.0936 | -106.7334
19.0973 | -106.7324
19.101 | -106.7302
19.1046 | -106.7282
19.1083 | -106.7269
19.112 | -106.7237
19.1157 | -106.7206
19.1193 | -106.7176
19.123 | -106.7134
19.1267 | -106.7056
19.1304 | -106.7016
19.134 | -106.6959
19.1377 | -106.691
19.1414 | -106.6852
19.145 | -106.6794
19.1487 | -106.675
19.1524 | -106.6702
19.1561 | -106.666
19.1597 | -106.6607
19.1634 | -106.6549
19.1671 | -106.6473
19.1708 | -106.6412
19.1744 | -106.6339
19.1781 | -106.6273
19.1818 | -106.6199
19.1855 | -106.6129
19.1891 | -106.6068
19.1928 | -106.6031
19.1965 | -106.6024
19.2001 | -106.5981
19.2038 | -106.5924
19.2075 | -106.5902
19.2112 | -106.5871
19.2148 | -106.5867
19.2185 | -106.5859
19.2222 | -106.5851
19.2259 | -106.5854
19.2295 | -106.5843
19.2332 | -106.5852
19.2369 | -106.5863
19.2406 | -106.5859
19.2442 | -106.5863
19.2479 | -106.5877
19.2516 | -106.5889
19.2552 | -106.5899
19.2589 | -106.5879
19.2626 | -106.5884
19.2663 | -106.591
19.2699 | -106.5911
19.2736 | -106.5911
19.2773 | -106.5918
19.281 | -106.5913
19.2846 | -106.5903
19.2883 | -106.5894
19.292 | -106.5882
19.2957 | -106.5865
19.2993 | -106.5833
19.303 | -106.5787
19.3067 | -106.5745
19.3103 | -106.5681
19.314 | -106.5623
19.3177 | -106.5558
19.3214 | -106.5496
19.325 | -106.5429
19.3287 | -106.5367
19.3324 | -106.5301
19.3361 | -106.523
19.3397 | -106.5162
19.3434 | -106.5077
19.3471 | -106.4983
19.3508 | -106.4897
19.3544 | -106.4804
19.3581 | -106.4704
19.3618 | -106.4618
19.3654 | -106.4547
19.3691 | -106.4475
19.3728 | -106.4403
19.3765 | -106.4342
19.3801 | -106.429
19.3838 | -106.4229
19.3875 | -106.4199
19.3912 | -106.4175
19.3948 | -106.4138
19.3985 | -106.4117
19.4022 | -106.4129
19.4059 | -106.4123
19.4095 | -106.4146
19.4132 | -106.4175
19.4169 | -106.4215
19.4205 | -106.4243
19.4242 | -106.43
19.4279 | -106.4374
19.4316 | -106.4435
19.4352 | -106.4504
19.4389 | -106.4595
19.4426 | -106.4689
19.4463 | -106.4789
19.4499 | -106.4898
19.4536 | -106.4995
19.4573 | -106.509
19.4609 | -106.5184
19.4646 | -106.5265
19.4683 | -106.5336
19.472 | -106.54
19.4756 | -106.5489
19.4793 | -106.5595
19.483 | -106.568
19.4867 | -106.5765
19.4903 | -106.5833
19.494 | -106.5899
19.4977 | -106.5954
19.5014 | -106.5996
19.505 | -106.6044
19.5087 | -106.6087
19.5124 | -106.6137
19.516 | -106.6175
19.5197 | -106.6212
19.5234 | -106.6246
19.5271 | -106.6261
19.5307 | -106.6282
19.5344 | -106.6314
19.5381 | -106.6325
19.5418 | -106.6292
19.5454 | -106.6258
19.5491 | -106.6236
19.5528 | -106.6194
19.5565 | -106.6161
19.5601 | -106.614
19.5638 | -106.6087
19.5675 | -106.6023
19.5711 | -106.5947
19.5748 | -106.5889
19.5785 | -106.5816
19.5822 | -106.5735
19.5858 | -106.5661
19.5895 | -106.5593
19.5932 | -106.5529
19.5969 | -106.5447
19.6005 | -106.536
19.6042 | -106.5279
19.6079 | -106.5207
19.6116 | -106.5147
19.6152 | -106.5088
19.6189 | -106.5028
19.6226 | -106.4983
19.6262 | -106.4948
19.6299 | -106.4901
19.6336 | -106.4863
19.6373 | -106.4831
19.6409 | -106.4776
19.6446 | -106.4742
19.6483 | -106.4744
19.652 | -106.4748
19.6556 | -106.4765
19.6593 | -106.478
19.663 | -106.4808
19.6667 | -106.4828
19.6703 | -106.4859
19.674 | -106.4929
19.6777 | -106.4992
19.6813 | -106.5047
19.685 | -106.5109
19.6887 | -106.5145
19.6924 | -106.518
19.696 | -106.5231
19.6997 | -106.5289
19.7034 | -106.5344
19.7071 | -106.5398
19.7107 | -106.5469
19.7144 | -106.5555
19.7181 | -106.5642
19.7218 | -106.5717
19.7254 | -106.5788
19.7291 | -106.5852
19.7328 | -106.5895
19.7364 | -106.5956
19.7401 | -106.6
19.7438 | -106.6068
19.7475 | -106.6102
19.7511 | -106.6123
19.7548 | -106.6146
19.7585 | -106.6131
19.7622 | -106.6117
19.7658 | -106.6112
19.7695 | -106.6106
19.7732 | -106.6083
19.7768 | -106.6088
19.7805 | -106.6071
19.7842 | -106.6042
19.7879 | -106.5995
19.7915 | -106.5937
19.7952 | -106.5862
19.7989 | -106.5779
19.8026 | -106.5704
19.8062 | -106.5641
19.8099 | -106.5556
19.8136 | -106.5477
19.8173 | -106.539
19.8209 | -106.5332
19.8246 | -106.5279
19.8283 | -106.5213
19.8319 | -106.5159
19.8356 | -106.5109
19.8393 | -106.5035
19.843 | -106.4977
19.8466 | -106.4937
19.8503 | -106.4895
19.854 | -106.4839
19.8577 | -106.4801
19.8613 | -106.4782
19.865 | -106.4767
19.8687 | -106.4745
19.8724 | -106.4721
19.876 | -106.4725
19.8797 | -106.4744
19.8834 | -106.4742
19.887 | -106.4765
19.8907 | -106.4782
19.8944 | -106.479
19.8981 | -106.4828
19.9017 | -106.4869
19.9054 | -106.4904
19.9091 | -106.4954
19.9128 | -106.5009
19.9164 | -106.5071
19.9201 | -106.5145
19.9238 | -106.5225
19.9275 | -106.5294
19.9311 | -106.5333
19.9348 | -106.5365
19.9385 | -106.5407
19.9421 | -106.5469
19.9458 | -106.5533
19.9495 | -106.5592
19.9532 | -106.5651
19.9568 | -106.5696
19.9605 | -106.5733
19.9642 | -106.5756
19.9679 | -106.5759
19.9715 | -106.5782
19.9752 | -106.5816
19.9789 | -106.5832
19.9826 | -106.585
19.9862 | -106.5866
19.9899 | -106.5859
19.9936 | -106.5838
19.9972 | -106.5818
20.0009 | -106.5789
20.0046 | -106.5768
20.0083 | -106.5739
20.0119 | -106.5712
20.0156 | -106.5685
20.0193 | -106.5634
20.023 | -106.5616
20.0266 | -106.5581
20.0303 | -106.5556
20.034 | -106.5559
20.0377 | -106.5501
20.0413 | -106.5487
20.045 | -106.5468
20.0487 | -106.543
20.0523 | -106.538
20.056 | -106.5336
20.0597 | -106.5281
20.0634 | -106.5229
20.067 | -106.5191
20.0707 | -106.5153
20.0744 | -106.5109
20.0781 | -106.508
20.0817 | -106.5036
20.0854 | -106.4981
20.0891 | -106.4912
20.0927 | -106.4878
20.0964 | -106.4826
20.1001 | -106.4761
20.1038 | -106.4695
20.1074 | -106.4613
20.1111 | -106.4526
20.1148 | -106.4423
20.1185 | -106.4297
20.1221 | -106.4195
20.1258 | -106.4096
20.1295 | -106.3985
20.1332 | -106.3882
20.1368 | -106.3689
20.1405 | -106.3538
20.1442 | -106.3386
20.1478 | -106.3216
20.1515 | -106.3053
20.1552 | -106.2894
20.1589 | -106.2733
20.1625 | -106.2531
20.1662 | -106.2342
20.1699 | -106.2178
20.1736 | -106.1998
20.1772 | -106.1802
20.1809 | -106.1613
20.1846 | -106.1405
20.1883 | -106.1205
20.1919 | -106.1025
20.1956 | -106.0854
20.1993 | -106.0676
20.2029 | -106.0485
20.2066 | -106.0297
20.2103 | -106.0123
20.214 | -105.9952
20.2176 | -105.9807
20.2213 | -105.9659
20.225 | -105.9506
20.2287 | -105.9357
20.2323 | -105.9218
20.236 | -105.909
20.2397 | -105.8968
20.2434 | -105.8868
20.247 | -105.8784
20.2507 | -105.8714
20.2544 | -105.8665
20.258 | -105.8605
20.2617 | -105.855
20.2654 | -105.8509
20.2691 | -105.8477
20.2727 | -105.8434
20.2764 | -105.8406
20.2801 | -105.8411
20.2838 | -105.8398
20.2874 | -105.8389
20.2911 | -105.841
20.2948 | -105.8451
20.2985 | -105.8479
20.3021 | -105.8511
20.3058 | -105.8574
20.3095 | -105.8609
20.3131 | -105.8655
20.3168 | -105.872
20.3205 | -105.8782
20.3242 | -105.8868
20.3278 | -105.8934
20.3315 | -105.9001
20.3352 | -105.9069
20.3389 | -105.9165
20.3425 | -105.9264
20.3462 | -105.9338
20.3499 | -105.9418
20.3536 | -105.9479
20.3572 | -105.9549
20.3609 | -105.961
20.3646 | -105.9677
20.3682 | -105.9739
20.3719 | -105.9777
20.3756 | -105.9821
20.3793 | -105.985
20.3829 | -105.9883
20.3866 | -105.99
20.3903 | -105.9899
20.394 | -105.9918
20.3976 | -105.992
20.4013 | -105.991
20.405 | -105.9904
20.4087 | -105.9881
20.4123 | -105.9844
20.416 | -105.9804
20.4197 | -105.9762
20.4233 | -105.9737
20.427 | -105.9691
20.4307 | -105.9652
20.4344 | -105.9598
20.438 | -105.9562
20.4417 | -105.9524
20.4454 | -105.9475
20.4491 | -105.9412
20.4527 | -105.9387
20.4564 | -105.9355
20.4601 | -105.9309
20.4637 | -105.9255
20.4674 | -105.9223
20.4711 | -105.9209
20.4748 | -105.919
20.4784 | -105.9159
20.4821 | -105.9125
20.4858 | -105.9096
20.4895 | -105.9079
20.4931 | -105.9079
20.4968 | -105.9073
20.5005 | -105.9083
20.5042 | -105.9104
20.5078 | -105.9128
20.5115 | -105.9173
20.5152 | -105.9216
20.5188 | -105.9253
20.5225 | -105.9295
20.5262 | -105.9352
20.5299 | -105.9398
20.5335 | -105.9439
20.5372 | -105.951
20.5409 | -105.9564
20.5446 | -105.9619
20.5482 | -105.968
20.5519 | -105.9742
20.5556 | -105.9787
20.5593 | -105.983
20.5629 | -105.9906
20.5666 | -105.9994
20.5703 | -106.0071
20.5739 | -106.0155
20.5776 | -106.0249
20.5813 | -106.032
20.585 | -106.0374
20.5886 | -106.0448
20.5923 | -106.0537
20.596 | -106.0607
20.5997 | -106.0671
20.6033 | -106.074
20.607 | -106.0777
20.6107 | -106.082
20.6144 | -106.089
20.618 | -106.0963
20.6217 | -106.1025
20.6254 | -106.1083
20.629 | -106.1117
20.6327 | -106.1156
20.6364 | -106.1218
20.6401 | -106.1276
20.6437 | -106.1305
20.6474 | -106.1345
20.6511 | -106.1386
20.6548 | -106.1391
20.6584 | -106.1399
20.6621 | -106.1409
20.6658 | -106.1366
20.6695 | -106.136
20.6731 | -106.1342
20.6768 | -106.1318
20.6805 | -106.1308
20.6841 | -106.1299
20.6878 | -106.1273
20.6915 | -106.1239
20.6952 | -106.1224
20.6988 | -106.1193
20.7025 | -106.115
20.7062 | -106.112
20.7099 | -106.1084
20.7135 | -106.104
20.7172 | -106.0981
20.7209 | -106.0923
20.7246 | -106.0877
20.7282 | -106.0847
20.7319 | -106.0799
20.7356 | -106.0744
20.7392 | -106.0699
20.7429 | -106.0677
20.7466 | -106.0637
20.7503 | -106.0587
20.7539 | -106.0537
20.7576 | -106.0529
20.7613 | -106.0517
20.765 | -106.0503
20.7686 | -106.0495
20.7723 | -106.0514
20.776 | -106.0538
20.7797 | -106.0559
20.7833 | -106.0602
20.787 | -106.0634
20.7907 | -106.0677
20.7943 | -106.0708
20.798 | -106.0746
20.8017 | -106.0795
20.8054 | -106.084
20.809 | -106.0895
20.8127 | -106.0953
20.8164 | -106.1026
20.8201 | -106.1087
20.8237 | -106.1158
20.8274 | -106.1227
20.8311 | -106.128
20.8347 | -106.1381
20.8384 | -106.1456
20.8421 | -106.1517
20.8458 | -106.1553
20.8494 | -106.1599
20.8531 | -106.1642
20.8568 | -106.1674
20.8605 | -106.17
20.8641 | -106.1719
20.8678 | -106.1745
20.8715 | -106.1765
20.8752 | -106.1759
20.8788 | -106.1751
20.8825 | -106.1753
20.8862 | -106.1754
20.8898 | -106.1706
20.8935 | -106.1665
20.8972 | -106.1629
20.9009 | -106.1622
20.9045 | -106.1577
20.9082 | -106.1497
20.9119 | -106.1433
20.9156 | -106.1384
20.9192 | -106.1336
20.9229 | -106.1296
20.9266 | -106.1248
20.9303 | -106.1204
20.9339 | -106.1175
20.9376 | -106.1157
20.9413 | -106.1126
20.9449 | -106.1102
20.9486 | -106.1099
20.9523 | -106.1087
20.956 | -106.1092
20.9596 | -106.1105
20.9633 | -106.1118
20.967 | -106.1163
20.9707 | -106.1215
20.9743 | -106.128
20.978 | -106.1351
20.9817 | -106.1421
20.9854 | -106.1493
20.989 | -106.1606
20.9927 | -106.1707
20.9964 | -106.1814
21 | -106.1937
21.0037 | -106.2069
21.0074 | -106.2187
21.0111 | -106.2321
21.0147 | -106.2488
21.0184 | -106.2666
21.0221 | -106.2803
21.0258 | -106.2922
21.0294 | -106.3054
21.0331 | -106.3207
21.0368 | -106.3374
21.0405 | -106.3544
21.0441 | -106.3681
21.0478 | -106.3829
21.0515 | -106.3984
21.0551 | -106.4108
21.0588 | -106.424
21.0625 | -106.4366
21.0662 | -106.4488
21.0698 | -106.4607
21.0735 | -106.47
21.0772 | -106.4811
21.0809 | -106.4898
21.0845 | -106.5004
21.0882 | -106.5101
21.0919 | -106.5203
21.0956 | -106.5289
21.0992 | -106.5352
21.1029 | -106.5404
21.1066 | -106.5454
21.1102 | -106.5518
21.1139 | -106.5569
21.1176 | -106.5627
21.1213 | -106.5665
21.1249 | -106.5691
21.1286 | -106.5722
21.1323 | -106.5734
21.136 | -106.5757
21.1396 | -106.579
21.1433 | -106.5834
21.147 | -106.5865
21.1506 | -106.5884
21.1543 | -106.5864
21.158 | -106.5843
21.1617 | -106.5828
21.1653 | -106.5817
21.169 | -106.5816
21.1727 | -106.5822
21.1764 | -106.5825
21.18 | -106.5808
21.1837 | -106.5782
21.1874 | -106.5781
21.1911 | -106.5777
21.1947 | -106.5741
21.1984 | -106.5713
21.2021 | -106.5681
21.2057 | -106.5656
21.2094 | -106.5639
21.2131 | -106.5601
21.2168 | -106.5568
21.2204 | -106.5512
21.2241 | -106.5432
21.2278 | -106.5355
21.2315 | -106.5286
21.2351 | -106.5215
21.2388 | -106.5146
21.2425 | -106.5056
21.2462 | -106.4975
21.2498 | -106.4899
21.2535 | -106.4819
21.2572 | -106.4724
21.2608 | -106.4628
21.2645 | -106.4534
21.2682 | -106.443
21.2719 | -106.4329
21.2755 | -106.4236
21.2792 | -106.4145
21.2829 | -106.4051
21.2866 | -106.396
21.2902 | -106.3857
21.2939 | -106.3762
21.2976 | -106.3681
21.3013 | -106.3592
21.3049 | -106.3484
21.3086 | -106.3406
21.3123 | -106.3323
21.3159 | -106.3227
21.3196 | -106.3144
21.3233 | -106.3101
21.327 | -106.307
21.3306 | -106.3021
21.3343 | -106.2963
21.338 | -106.2907
21.3417 | -106.2857
21.3453 | -106.2868
21.349 | -106.2875
21.3527 | -106.2887
21.3564 | -106.2899
21.36 | -106.2911
21.3637 | -106.2918
21.3674 | -106.2952
21.371 | -106.2992
21.3747 | -106.3052
21.3784 | -106.3086
21.3821 | -106.3132
21.3857 | -106.316
21.3894 | -106.3212
21.3931 | -106.3247
21.3968 | -106.3315
21.4004 | -106.3368
21.4041 | -106.3433
21.4078 | -106.3507
21.4115 | -106.3569
21.4151 | -106.3602
21.4188 | -106.3627
21.4225 | -106.3697
21.4261 | -106.3752
21.4298 | -106.3788
21.4335 | -106.382
21.4372 | -106.3846
21.4408 | -106.3868
21.4445 | -106.389
21.4482 | -106.3919
21.4519 | -106.3949
21.4555 | -106.399
21.4592 | -106.4008
21.4629 | -106.4022
21.4666 | -106.4021
21.4702 | -106.3998
21.4739 | -106.3987
21.4776 | -106.3976
21.4812 | -106.3979
21.4849 | -106.3959
21.4886 | -106.3951
21.4923 | -106.3953
21.4959 | -106.3941
21.4996 | -106.3947
21.5033 | -106.3934
21.507 | -106.3925
21.5106 | -106.3923
21.5143 | -106.3908
21.518 | -106.39
21.5216 | -106.389
21.5253 | -106.3922
21.529 | -106.3953
21.5327 | -106.3974
21.5363 | -106.3978
21.54 | -106.398
21.5437 | -106.4018
21.5474 | -106.4057
21.551 | -106.4066
21.5547 | -106.409
21.5584 | -106.4131
21.5621 | -106.4166
21.5657 | -106.4194
21.5694 | -106.4225
21.5731 | -106.4254
21.5767 | -106.4286
21.5804 | -106.4327
21.5841 | -106.4368
21.5878 | -106.4409
21.5914 | -106.4461
21.5951 | -106.4497
21.5988 | -106.4539
21.6025 | -106.4577
21.6061 | -106.4622
21.6098 | -106.4681
21.6135 | -106.4729
21.6172 | -106.4773
21.6208 | -106.4837
21.6245 | -106.4862
21.6282 | -106.4884
21.6318 | -106.4898
21.6355 | -106.4922
21.6392 | -106.4945
21.6429 | -106.4943
21.6465 | -106.4953
21.6502 | -106.4937
21.6539 | -106.4919
21.6576 | -106.4919
21.6612 | -106.4922
21.6649 | -106.4922
21.6686 | -106.4934
21.6723 | -106.492
21.6759 | -106.4898
21.6796 | -106.4887
21.6833 | -106.4877
21.6869 | -106.4885
21.6906 | -106.4903
21.6943 | -106.4895
21.698 | -106.4888
21.7016 | -106.4877
21.7053 | -106.4885
21.709 | -106.488
21.7127 | -106.4848
21.7163 | -106.4801
21.72 | -106.4762
21.7237 | -106.4726
21.7274 | -106.4689
21.731 | -106.4631
21.7347 | -106.4574
21.7384 | -106.4515
21.742 | -106.4449
21.7457 | -106.439
21.7494 | -106.4319
21.7531 | -106.4241
21.7567 | -106.4153
21.7604 | -106.4051
21.7641 | -106.3962
21.7678 | -106.3864
21.7714 | -106.3775
21.7751 | -106.3678
21.7788 | -106.3571
21.7825 | -106.3467
21.7861 | -106.3369
21.7898 | -106.3241
21.7935 | -106.3115
21.7971 | -106.2979
21.8008 | -106.2825
21.8045 | -106.2661
21.8082 | -106.2498
21.8118 | -106.2336
21.8155 | -106.2152
21.8192 | -106.1971
21.8229 | -106.1806
21.8265 | -106.1603
21.8302 | -106.1375
21.8339 | -106.1159
21.8376 | -106.0966
21.8412 | -106.0765
21.8449 | -106.0562
21.8486 | -106.0366
21.8522 | -106.0139
21.8559 | -105.9938
21.8596 | -105.9748
21.8633 | -105.954
21.8669 | -105.9334
21.8706 | -105.9128
21.8743 | -105.8926
21.878 | -105.8729
21.8816 | -105.8524
21.8853 | -105.8318
21.889 | -105.8093
21.8926 | -105.7907
21.8963 | -105.7745
21.9 | -105.759
21.9037 | -105.7436
21.9073 | -105.7272
21.911 | -105.7122
21.9147 | -105.6976
21.9184 | -105.684
21.922 | -105.6716
21.9257 | -105.6586
21.9294 | -105.6449
21.9331 | -105.6345
21.9367 | -105.6254
21.9404 | -105.6172
21.9441 | -105.6105
21.9477 | -105.6049
21.9514 | -105.6014
21.9551 | -105.5976
21.9588 | -105.5944
21.9624 | -105.5925
21.9661 | -105.5894
21.9698 | -105.5876
21.9735 | -105.5891
21.9771 | -105.5916
21.9808 | -105.5954
21.9845 | -105.5996
21.9882 | -105.6047
21.9918 | -105.6117
21.9955 | -105.6182
21.9992 | -105.627
22.0028 | -105.6364
22.0065 | -105.649
22.0102 | -105.6612
22.0139 | -105.6744
22.0175 | -105.6889
22.0212 | -105.7041
22.0249 | -105.7214
22.0286 | -105.7417
22.0322 | -105.7629
22.0359 | -105.7824
22.0396 | -105.8016
22.0433 | -105.8232
22.0469 | -105.8474
22.0506 | -105.8706
22.0543 | -105.8932
22.0579 | -105.9171
22.0616 | -105.9407
22.0653 | -105.9649
22.069 | -105.9889
22.0726 | -106.0142
22.0763 | -106.0385
22.08 | -106.063
22.0837 | -106.0857
22.0873 | -106.1106
22.091 | -106.1341
22.0947 | -106.1565
22.0984 | -106.1771
22.102 | -106.1985
22.1057 | -106.2196
22.1094 | -106.2407
22.113 | -106.2607
22.1167 | -106.2811
22.1204 | -106.2992
22.1241 | -106.3168
22.1277 | -106.3343
22.1314 | -106.3476
22.1351 | -106.359
22.1388 | -106.3723
22.1424 | -106.3831
22.1461 | -106.3906
22.1498 | -106.3959
22.1535 | -106.4021
22.1571 | -106.4058
22.1608 | -106.4091
22.1645 | -106.4114
22.1681 | -106.4096
22.1718 | -106.4053
22.1755 | -106.4001
22.1792 | -106.393
22.1828 | -106.3883
22.1865 | -106.3823
22.1902 | -106.3733
22.1939 | -106.364
22.1975 | -106.3555
22.2012 | -106.3453
22.2049 | -106.3316
22.2085 | -106.315
22.2122 | -106.2991
22.2159 | -106.2836
22.2196 | -106.2682
22.2232 | -106.2505
22.2269 | -106.2318
22.2306 | -106.2132
22.2343 | -106.1929
22.2379 | -106.1747
22.2416 | -106.158
22.2453 | -106.1416
22.249 | -106.1219
22.2526 | -106.1028
22.2563 | -106.0856
22.26 | -106.0667
22.2636 | -106.049
22.2673 | -106.0315
22.271 | -106.0149
22.2747 | -106.001
22.2783 | -105.9867
22.282 | -105.9731
22.2857 | -105.9599
22.2894 | -105.9452
22.293 | -105.9303
22.2967 | -105.918
22.3004 | -105.9084
22.3041 | -105.8987
22.3077 | -105.8885
22.3114 | -105.8793
22.3151 | -105.8706
22.3187 | -105.8653
22.3224 | -105.8582
22.3261 | -105.8524
22.3298 | -105.8468
22.3334 | -105.8445
22.3371 | -105.8419
22.3408 | -105.8371
22.3445 | -105.8328
22.3481 | -105.8298
22.3518 | -105.8283
22.3555 | -105.8285
22.3592 | -105.829
22.3628 | -105.8296
22.3665 | -105.8324
22.3702 | -105.8359
22.3738 | -105.8381
22.3775 | -105.8418
22.3812 | -105.8458
22.3849 | -105.8492
22.3885 | -105.8533
22.3922 | -105.8584
22.3959 | -105.8639
22.3996 | -105.868
22.4032 | -105.8722
22.4069 | -105.8792
22.4106 | -105.8852
22.4143 | -105.8897
22.4179 | -105.8946
22.4216 | -105.8999
22.4253 | -105.9033
22.4289 | -105.8996
22.4326 | -105.8964
22.4363 | -105.8833
22.44 | -105.8559
22.4436 | -105.8279
22.4473 | -105.8095
22.451 | -105.7951
22.4547 | -105.7816
22.4583 | -105.7683
22.462 | -105.7518
22.4657 | -105.7329
22.4694 | -105.7105
22.473 | -105.6874
22.4767 | -105.6599
22.4804 | -105.6319
22.484 | -105.6029
22.4877 | -105.5722
22.4914 | -105.5392
22.4951 | -105.5045
22.4987 | -105.4665
22.5024 | -105.4254
22.5061 | -105.3829
22.5098 | -105.34
22.5134 | -105.2954
22.5171 | -105.2514
22.5208 | -105.2059
22.5245 | -105.1617
22.5281 | -105.1161
22.5318 | -105.0702
22.5355 | -105.0248
22.5391 | -104.9814
22.5428 | -104.9396
22.5465 | -104.8968
22.5502 | -104.8575
22.5538 | -104.8204
22.5575 | -104.7848
22.5612 | -104.751
22.5649 | -104.7215
22.5685 | -104.6943
22.5722 | -104.6699
22.5759 | -104.6493
22.5795 | -104.6322
22.5832 | -104.6179
22.5869 | -104.6075
22.5906 | -104.6015
22.5942 | -104.5989
22.5979 | -104.6013
22.6016 | -104.6082
22.6053 | -104.6183
22.6089 | -104.6334
22.6126 | -104.6538
22.6163 | -104.6763
22.62 | -104.7014
22.6236 | -104.7305
22.6273 | -104.7641
22.631 | -104.8007
22.6346 | -104.8396
22.6383 | -104.8807
22.642 | -104.9261
22.6457 | -104.9735
22.6493 | -105.0249
22.653 | -105.0764
22.6567 | -105.1308
22.6604 | -105.1877
22.664 | -105.245
22.6677 | -105.3019
22.6714 | -105.3582
22.6751 | -105.4143
22.6787 | -105.47
22.6824 | -105.5259
22.6861 | -105.5817
22.6897 | -105.6356
22.6934 | -105.688
22.6971 | -105.7401
22.7008 | -105.7908
22.7044 | -105.8395
22.7081 | -105.8864
22.7118 | -105.9314
22.7155 | -105.9754
22.7191 | -106.0179
22.7228 | -106.0567
22.7265 | -106.0928
22.7302 | -106.1262
22.7338 | -106.1593
22.7375 | -106.1894
22.7412 | -106.2164
22.7448 | -106.2431
22.7485 | -106.2739
22.7522 | -106.2998
22.7559 | -106.3337
22.7595 | -106.3753
22.7632 | -106.4091
22.7669 | -106.4279
22.7706 | -106.4369
22.7742 | -106.4393
22.7779 | -106.4404
22.7816 | -106.4421
22.7853 | -106.4445
22.7889 | -106.4456
22.7926 | -106.4495
22.7963 | -106.4528
22.7999 | -106.4572
22.8036 | -106.46
22.8073 | -106.4619
22.811 | -106.466
22.8146 | -106.4691
22.8183 | -106.4712
22.822 | -106.4726
22.8257 | -106.4755
22.8293 | -106.4809
22.833 | -106.4834
22.8367 | -106.4857
22.8404 | -106.489
22.844 | -106.4912
22.8477 | -106.4915
22.8514 | -106.4916
22.855 | -106.4944
22.8587 | -106.4976
22.8624 | -106.4992
22.8661 | -106.5015
22.8697 | -106.5048
22.8734 | -106.5066
22.8771 | -106.5073
22.8808 | -106.5073
22.8844 | -106.5067
22.8881 | -106.5087
22.8918 | -106.5114
22.8955 | -106.5142
22.8991 | -106.5161
22.9028 | -106.5178
22.9065 | -106.5198
22.9101 | -106.5224
22.9138 | -106.5239
22.9175 | -106.5245
22.9212 | -106.5244
22.9248 | -106.5253
22.9285 | -106.5262
22.9322 | -106.5257
22.9359 | -106.5266
22.9395 | -106.5297
22.9432 | -106.5314
22.9469 | -106.5319
22.9505 | -106.5315
22.9542 | -106.5324
22.9579 | -106.5351
22.9616 | -106.537
22.9652 | -106.5371
22.9689 | -106.5378
22.9726 | -106.5381
22.9763 | -106.5405
22.9799 | -106.5416
22.9836 | -106.5421
22.9873 | -106.5414
22.991 | -106.5411
22.9946 | -106.5433
22.9983 | -106.5454
23.002 | -106.5469
23.0056 | -106.5486
23.0093 | -106.5497
23.013 | -106.5503
23.0167 | -106.5519
23.0203 | -106.5519
23.024 | -106.5527
23.0277 | -106.5541
23.0314 | -106.5525
23.035 | -106.5505
23.0387 | -106.5478
23.0424 | -106.5451
23.0461 | -106.5446
23.0497 | -106.5425
23.0534 | -106.5405
23.0571 | -106.5383
23.0607 | -106.5363
23.0644 | -106.5341
23.0681 | -106.5298
23.0718 | -106.526
23.0754 | -106.5232
23.0791 | -106.5202
23.0828 | -106.5181
23.0865 | -106.5142
23.0901 | -106.51
23.0938 | -106.5074
23.0975 | -106.5079
23.1012 | -106.5073
23.1048 | -106.5042
23.1085 | -106.5009
23.1122 | -106.4981
23.1158 | -106.4976
23.1195 | -106.4946
23.1232 | -106.4922
23.1269 | -106.4899
23.1305 | -106.4893
23.1342 | -106.4893
23.1379 | -106.4873
23.1416 | -106.487
23.1452 | -106.4865
23.1489 | -106.4874
23.1526 | -106.4908
23.1563 | -106.4916
23.1599 | -106.4933
23.1636 | -106.4953
23.1673 | -106.4954
23.1709 | -106.4984
23.1746 | -106.5023
23.1783 | -106.5046
23.182 | -106.5068
23.1856 | -106.5086
23.1893 | -106.5116
23.193 | -106.5107
23.1967 | -106.5121
23.2003 | -106.5142
23.204 | -106.5152
23.2077 | -106.5168
23.2113 | -106.521
23.215 | -106.5211
23.2187 | -106.5203
23.2224 | -106.5212
23.226 | -106.521
23.2297 | -106.5199
23.2334 | -106.52
23.2371 | -106.5199
23.2407 | -106.5175
23.2444 | -106.5152
23.2481 | -106.5126
23.2518 | -106.5087
23.2554 | -106.5067
23.2591 | -106.504
23.2628 | -106.499
23.2664 | -106.4945
23.2701 | -106.4921
23.2738 | -106.4901
23.2775 | -106.4871
23.2811 | -106.4824
23.2848 | -106.476
23.2885 | -106.4708
23.2922 | -106.4657
23.2958 | -106.4603
23.2995 | -106.455
23.3032 | -106.4483
23.3069 | -106.4424
23.3105 | -106.4359
23.3142 | -106.4307
23.3179 | -106.4267
23.3215 | -106.4239
23.3252 | -106.4194
23.3289 | -106.4148
23.3326 | -106.412
23.3362 | -106.4076
23.3399 | -106.4038
23.3436 | -106.402
23.3473 | -106.3998
23.3509 | -106.398
23.3546 | -106.3959
23.3583 | -106.3946
23.362 | -106.394
23.3656 | -106.393
23.3693 | -106.3932
23.373 | -106.3948
23.3766 | -106.396
23.3803 | -106.3989
23.384 | -106.4021
23.3877 | -106.4045
23.3913 | -106.4083
23.395 | -106.4136
23.3987 | -106.4191
23.4024 | -106.4252
23.406 | -106.4308
23.4097 | -106.4357
23.4134 | -106.4405
23.4171 | -106.447
23.4207 | -106.4535
23.4244 | -106.4597
23.4281 | -106.4658
23.4317 | -106.4733
23.4354 | -106.4829
23.4391 | -106.4927
23.4428 | -106.4993
23.4464 | -106.5045
23.4501 | -106.5106
23.4538 | -106.517
23.4575 | -106.5226
23.4611 | -106.5259
23.4648 | -106.5298
23.4685 | -106.5354
23.4722 | -106.5388
23.4758 | -106.5426
23.4795 | -106.5467
23.4832 | -106.5475
23.4868 | -106.5487
23.4905 | -106.5502
23.4942 | -106.5499
23.4979 | -106.5497
23.5015 | -106.5499
23.5052 | -106.5479
23.5089 | -106.5456
23.5126 | -106.5415
23.5162 | -106.5375
23.5199 | -106.5323
23.5236 | -106.5271
23.5273 | -106.5212
23.5309 | -106.514
23.5346 | -106.5063
23.5383 | -106.4975
23.5419 | -106.487
23.5456 | -106.4787
23.5493 | -106.4702
23.553 | -106.459
23.5566 | -106.4481
23.5603 | -106.4372
23.564 | -106.4263
23.5677 | -106.4135
23.5713 | -106.4023
23.575 | -106.3915
23.5787 | -106.3821
23.5824 | -106.3733
23.586 | -106.3634
23.5897 | -106.3536
23.5934 | -106.3451
23.597 | -106.3363
23.6007 | -106.3266
23.6044 | -106.3184
23.6081 | -106.312
23.6117 | -106.3064
23.6154 | -106.3026
23.6191 | -106.298
23.6228 | -106.2944
23.6264 | -106.2913
23.6301 | -106.2894
23.6338 | -106.2899
23.6374 | -106.2916
23.6411 | -106.2945
23.6448 | -106.297
23.6485 | -106.2995
23.6521 | -106.3014
23.6558 | -106.3024
23.6595 | -106.3061
23.6632 | -106.3128
23.6668 | -106.3196
23.6705 | -106.3262
23.6742 | -106.3345
23.6779 | -106.3437
23.6815 | -106.3537
23.6852 | -106.3641
23.6889 | -106.3745
23.6925 | -106.3848
23.6962 | -106.3962
23.6999 | -106.4088
23.7036 | -106.4207
23.7072 | -106.4323
23.7109 | -106.4442
23.7146 | -106.4565
23.7183 | -106.4698
23.7219 | -106.4827
23.7256 | -106.4944
23.7293 | -106.5047
23.733 | -106.5151
23.7366 | -106.5253
23.7403 | -106.5356
23.744 | -106.5458
23.7476 | -106.5563
23.7513 | -106.5673
23.755 | -106.5763
23.7587 | -106.5858
23.7623 | -106.5943
23.766 | -106.603
23.7697 | -106.6122
23.7734 | -106.6201
23.777 | -106.627
23.7807 | -106.6343
23.7844 | -106.6418
23.7881 | -106.6469
23.7917 | -106.6525
23.7954 | -106.6584
23.7991 | -106.662
23.8027 | -106.6639
23.8064 | -106.6656
23.8101 | -106.6665
23.8138 | -106.6666
23.8174 | -106.6671
23.8211 | -106.6681
23.8248 | -106.6692
23.8285 | -106.6707
23.8321 | -106.6707
23.8358 | -106.6672
23.8395 | -106.6659
23.8432 | -106.6662
23.8468 | -106.6657
23.8505 | -106.6654
23.8542 | -106.6648
23.8578 | -106.6643
23.8615 | -106.6648
23.8652 | -106.6643
23.8689 | -106.6644
23.8725 | -106.665
23.8762 | -106.6657
23.8799 | -106.6677
23.8836 | -106.6673
23.8872 | -106.6673
23.8909 | -106.668
23.8946 | -106.6692
23.8983 | -106.6701
23.9019 | -106.6688
23.9056 | -106.6703
23.9093 | -106.6718
23.9129 | -106.673
23.9166 | -106.6747
23.9203 | -106.6755
23.924 | -106.6759
23.9276 | -106.6787
23.9313 | -106.68
23.935 | -106.6811
23.9387 | -106.6836
23.9423 | -106.6854
23.946 | -106.6874
23.9497 | -106.6924
23.9534 | -106.696
23.957 | -106.6994
23.9607 | -106.704
23.9644 | -106.7083
23.968 | -106.7136
23.9717 | -106.7167
23.9754 | -106.7196
23.9791 | -106.7239
23.9827 | -106.7269
23.9864 | -106.732
23.9901 | -106.7381
23.9938 | -106.7441
23.9974 | -106.7489
24.0011 | -106.751
24.0048 | -106.7531
24.0084 | -106.7578
24.0121 | -106.762
24.0158 | -106.7644
24.0195 | -106.7666
24.0231 | -106.769
24.0268 | -106.7709
24.0305 | -106.7751
24.0342 | -106.7792
24.0378 | -106.7823
24.0415 | -106.7819
24.0452 | -106.7819
24.0489 | -106.7818
24.0525 | -106.7818
24.0562 | -106.7811
24.0599 | -106.7816
24.0635 | -106.7813
24.0672 | -106.7825
24.0709 | -106.7848
24.0746 | -106.7859
24.0782 | -106.7866
24.0819 | -106.7878
24.0856 | -106.7872
24.0893 | -106.7857
24.0929 | -106.7838
24.0966 | -106.7808
24.1003 | -106.7789
24.104 | -106.7774
24.1076 | -106.7748
24.1113 | -106.7732
24.115 | -106.7737
24.1186 | -106.7739
24.1223 | -106.7725
24.126 | -106.7711
24.1297 | -106.7705
24.1333 | -106.771
24.137 | -106.7706
24.1407 | -106.7699
24.1444 | -106.7701
24.148 | -106.7681
24.1517 | -106.7667
24.1554 | -106.7667
24.1591 | -106.7663
24.1627 | -106.7669
24.1664 | -106.7687
24.1701 | -106.7717
24.1737 | -106.7743
24.1774 | -106.7768
24.1811 | -106.7787
24.1848 | -106.7815
24.1884 | -106.7863
24.1921 | -106.7902
24.1958 | -106.7943
24.1995 | -106.7988
24.2031 | -106.8033
24.2068 | -106.8055
24.2105 | -106.8088
24.2142 | -106.8139
24.2178 | -106.8184
24.2215 | -106.8226
24.2252 | -106.8287
24.2288 | -106.8349
24.2325 | -106.8397
24.2362 | -106.845
24.2399 | -106.8506
24.2435 | -106.8566
24.2472 | -106.8594
24.2509 | -106.8624
24.2546 | -106.8661
24.2582 | -106.8662
24.2619 | -106.8665
24.2656 | -106.8681
24.2692 | -106.8643
24.2729 | -106.8513
24.2766 | -106.8268
24.2803 | -106.7713
24.2839 | -106.6857
24.2876 | -106.6134
24.2913 | -106.5527
24.295 | -106.5063
24.2986 | -106.4633
24.3023 | -106.4173
24.306 | -106.3685
24.3097 | -106.313
24.3133 | -106.2559
24.317 | -106.1962
24.3207 | -106.1324
24.3243 | -106.0665
24.328 | -105.9955
24.3317 | -105.922
24.3354 | -105.8468
24.339 | -105.7681
24.3427 | -105.6876
24.3464 | -105.6068
24.3501 | -105.5229
24.3537 | -105.4377
24.3574 | -105.3523
24.3611 | -105.265
24.3648 | -105.1769
24.3684 | -105.0894
24.3721 | -105.002
24.3758 | -104.9135
24.3794 | -104.8335
24.3831 | -104.7545
24.3868 | -104.6779
24.3905 | -104.6026
24.3941 | -104.5286
24.3978 | -104.4573
24.4015 | -104.3909
24.4052 | -104.3276
24.4088 | -104.2678
24.4125 | -104.2122
24.4162 | -104.1602
24.4199 | -104.1118
24.4235 | -104.0678
24.4272 | -104.0297
24.4309 | -103.9973
24.4345 | -103.969
24.4382 | -103.9466
24.4419 | -103.9292
24.4456 | -103.9153
24.4492 | -103.9065
24.4529 | -103.9032
24.4566 | -103.905
24.4603 | -103.9122
24.4639 | -103.9244
24.4676 | -103.9412
24.4713 | -103.964
24.475 | -103.991
24.4786 | -104.0219
24.4823 | -104.058
24.486 | -104.0987
24.4896 | -104.143
24.4933 | -104.1927
24.497 | -104.2449
24.5007 | -104.2997
24.5043 | -104.3582
24.508 | -104.4187
24.5117 | -104.4815
24.5154 | -104.5468
24.519 | -104.6145
24.5227 | -104.6833
24.5264 | -104.7539
24.5301 | -104.8242
24.5337 | -104.8953
24.5374 | -104.968
24.5411 | -105.0385
24.5447 | -105.1085
24.5484 | -105.1788
24.5521 | -105.2496
24.5558 | -105.3179
24.5594 | -105.3833
24.5631 | -105.4462
24.5668 | -105.5052
24.5705 | -105.5571
24.5741 | -105.5946
24.5778 | -105.6074
24.5815 | -105.559
24.5852 | -105.4223
24.5888 | -105.3041
24.5925 | -105.231
24.5962 | -105.1699
24.5998 | -105.1157
24.6035 | -105.0564
24.6072 | -104.9879
24.6109 | -104.9104
24.6145 | -104.8228
24.6182 | -104.7293
24.6219 | -104.6332
24.6256 | -104.5446
24.6292 | -104.4604
24.6329 | -104.3539
24.6366 | -104.2293
24.6403 | -104.0898
24.6439 | -103.9439
24.6476 | -103.7962
24.6513 | -103.6475
24.6549 | -103.4999
24.6586 | -103.3522
24.6623 | -103.2043
24.666 | -103.0581
24.6696 | -102.9137
24.6733 | -102.7716
24.677 | -102.6315
24.6807 | -102.494
24.6843 | -102.3602
24.688 | -102.2301
24.6917 | -102.1037
24.6953 | -101.9832
24.699 | -101.8678
24.7027 | -101.7573
24.7064 | -101.6533
24.71 | -101.5554
24.7137 | -101.4631
24.7174 | -101.3766
24.7211 | -101.2968
24.7247 | -101.2237
24.7284 | -101.1565
24.7321 | -101.0968
24.7358 | -101.0449
24.7394 | -101.0005
24.7431 | -100.9634
24.7468 | -100.9334
24.7504 | -100.9104
24.7541 | -100.895
24.7578 | -100.8877
24.7615 | -100.8874
24.7651 | -100.8947
24.7688 | -100.9095
24.7725 | -100.9319
24.7762 | -100.962
24.7798 | -100.9997
24.7835 | -101.045
24.7872 | -101.0972
24.7909 | -101.1564
24.7945 | -101.2228
24.7982 | -101.2956
24.8019 | -101.375
24.8055 | -101.4607
24.8092 | -101.5528
24.8129 | -101.6515
24.8166 | -101.7561
24.8202 | -101.8658
24.8239 | -101.9807
24.8276 | -102.1006
24.8313 | -102.2249
24.8349 | -102.3527
24.8386 | -102.4845
24.8423 | -102.6194
24.846 | -102.7574
24.8496 | -102.8974
24.8533 | -103.0396
24.857 | -103.1836
24.8606 | -103.3285
24.8643 | -103.473
24.868 | -103.6176
24.8717 | -103.7626
24.8753 | -103.9058
24.879 | -104.0466
24.8827 | -104.1843
24.8864 | -104.3195
24.89 | -104.4536
24.8937 | -104.5831
24.8974 | -104.7075
24.9011 | -104.8266
24.9047 | -104.8759
24.9084 | -104.987
24.9121 | -105.0933
24.9157 | -105.1954
24.9194 | -105.2919
24.9231 | -105.3856
24.9268 | -105.4835
24.9304 | -105.5924
24.9341 | -105.7397
24.9378 | -105.9484
24.9415 | -106.1037
24.9451 | -106.1792
24.9488 | -106.219
24.9525 | -106.2332
24.9562 | -106.2417
24.9598 | -106.2499
24.9635 | -106.2566
24.9672 | -106.2662
24.9708 | -106.2773
24.9745 | -106.2868
24.9782 | -106.2916
24.9819 | -106.2914
24.9855 | -106.2903
24.9892 | -106.2886
24.9929 | -106.2867
24.9966 | -106.2841
25.0002 | -106.2823
25.0039 | -106.2795
25.0076 | -106.2746
25.0112 | -106.2719
25.0149 | -106.2693
25.0186 | -106.2641
25.0223 | -106.2599
25.0259 | -106.2547
25.0296 | -106.251
25.0333 | -106.2479
25.037 | -106.2438
25.0406 | -106.241
25.0443 | -106.2394
25.048 | -106.2357
25.0517 | -106.2318
25.0553 | -106.229
25.059 | -106.2239
25.0627 | -106.2197
25.0663 | -106.216
25.07 | -106.2122
25.0737 | -106.2091
25.0774 | -106.206
25.081 | -106.2042
25.0847 | -106.2015
25.0884 | -106.1986
25.0921 | -106.1962
25.0957 | -106.1942
25.0994 | -106.1928
25.1031 | -106.1894
25.1068 | -106.1849
25.1104 | -106.1808
25.1141 | -106.1763
25.1178 | -106.1729
25.1214 | -106.1713
25.1251 | -106.1694
25.1288 | -106.1654
25.1325 | -106.1607
25.1361 | -106.1571
25.1398 | -106.1524
25.1435 | -106.1479
25.1472 | -106.1447
25.1508 | -106.1427
25.1545 | -106.1381
25.1582 | -106.1328
25.1619 | -106.1259
25.1655 | -106.1183
25.1692 | -106.1106
25.1729 | -106.1038
25.1765 | -106.0969
25.1802 | -106.0902
25.1839 | -106.0827
25.1876 | -106.0762
25.1912 | -106.0683
25.1949 | -106.0578
25.1986 | -106.047
25.2023 | -106.0388
25.2059 | -106.03
25.2096 | -106.0191
25.2133 | -106.0077
25.217 | -105.9987
25.2206 | -105.9882
25.2243 | -105.9771
25.228 | -105.9672
25.2316 | -105.9579
25.2353 | -105.9473
25.239 | -105.9365
25.2427 | -105.9242
25.2463 | -105.9106
25.25 | -105.8999
25.2537 | -105.8897
25.2574 | -105.8786
25.261 | -105.867
25.2647 | -105.8552
25.2684 | -105.8427
25.2721 | -105.8295
25.2757 | -105.8189
25.2794 | -105.8087
25.2831 | -105.7975
25.2867 | -105.7845
25.2904 | -105.7726
25.2941 | -105.7629
25.2978 | -105.7514
25.3014 | -105.7397
25.3051 | -105.7279
25.3088 | -105.717
25.3125 | -105.7075
25.3161 | -105.6981
25.3198 | -105.6888
25.3235 | -105.6787
25.3271 | -105.668
25.3308 | -105.6578
25.3345 | -105.6488
25.3382 | -105.64
25.3418 | -105.6323
25.3455 | -105.6243
25.3492 | -105.6159
25.3529 | -105.609
25.3565 | -105.603
25.3602 | -105.5965
25.3639 | -105.5903
25.3676 | -105.5827
25.3712 | -105.5779
25.3749 | -105.5748
25.3786 | -105.5705
25.3822 | -105.5674
25.3859 | -105.5639
25.3896 | -105.5608
25.3933 | -105.5573
25.3969 | -105.5533
25.4006 | -105.5518
25.4043 | -105.5506
25.408 | -105.5508
25.4116 | -105.5515
25.4153 | -105.5536
25.419 | -105.5563
25.4227 | -105.5596
25.4263 | -105.5637
25.43 | -105.5679
25.4337 | -105.5783
25.4373 | -105.5821
25.441 | -105.5862
25.4447 | -105.5903
25.4484 | -105.5948
25.452 | -105.5999
25.4557 | -105.6066
25.4594 | -105.6148
25.4631 | -105.624
25.4667 | -105.6329
25.4704 | -105.6417
25.4741 | -105.6504
25.4778 | -105.659
25.4814 | -105.6682
25.4851 | -105.6756
25.4888 | -105.6827
25.4924 | -105.6907
25.4961 | -105.6987
25.4998 | -105.7091
25.5035 | -105.718
25.5071 | -105.7272
25.5108 | -105.7351
25.5145 | -105.7453
25.5182 | -105.7566
25.5218 | -105.767
25.5255 | -105.7766
25.5292 | -105.786
25.5329 | -105.7937
25.5365 | -105.8018
25.5402 | -105.8091
25.5439 | -105.8152
25.5475 | -105.8213
25.5512 | -105.8278
25.5549 | -105.8347
25.5586 | -105.84
25.5622 | -105.8429
25.5659 | -105.8456
25.5696 | -105.85
25.5733 | -105.8548
25.5769 | -105.8593
25.5806 | -105.863
25.5843 | -105.8648
25.588 | -105.8651
25.5916 | -105.8672
25.5953 | -105.8694
25.599 | -105.8706
25.6026 | -105.8732
25.6063 | -105.8771
25.61 | -105.8772
25.6137 | -105.8755
25.6173 | -105.8746
25.621 | -105.8752
25.6247 | -105.8755
25.6284 | -105.8764
25.632 | -105.8779
25.6357 | -105.8777
25.6394 | -105.8774
25.6431 | -105.8777
25.6467 | -105.8795
25.6504 | -105.8801
25.6541 | -105.8781
25.6577 | -105.8775
25.6614 | -105.8775
25.6651 | -105.8772
25.6688 | -105.8775
25.6724 | -105.8763
25.6761 | -105.8763
25.6798 | -105.8771
25.6835 | -105.879
25.6871 | -105.8796
25.6908 | -105.8799
25.6945 | -105.8799
25.6981 | -105.8793
25.7018 | -105.8788
25.7055 | -105.8792
25.7092 | -105.882
25.7128 | -105.8845
25.7165 | -105.8859
25.7202 | -105.8891
25.7239 | -105.8931
25.7275 | -105.8969
25.7312 | -105.9
25.7349 | -105.9035
25.7386 | -105.9077
25.7422 | -105.911
25.7459 | -105.9152
25.7496 | -105.9194
25.7532 | -105.9243
25.7569 | -105.9305
25.7606 | -105.9323
25.7643 | -105.9344
25.7679 | -105.9376
25.7716 | -105.9415
25.7753 | -105.9447
25.779 | -105.9464
25.7826 | -105.9504
25.7863 | -105.9549
25.79 | -105.9572
25.7937 | -105.9587
25.7973 | -105.9626
25.801 | -105.9679
25.8047 | -105.9713
25.8083 | -105.9741
25.812 | -105.9769
25.8157 | -105.9785
25.8194 | -105.9796
25.823 | -105.9795
25.8267 | -105.9804
25.8304 | -105.9809
25.8341 | -105.9804
25.8377 | -105.9813
25.8414 | -105.9796
25.8451 | -105.9773
25.8488 | -105.9772
25.8524 | -105.9768
25.8561 | -105.9752
25.8598 | -105.9733
25.8634 | -105.9709
25.8671 | -105.9708
25.8708 | -105.9703
25.8745 | -105.9701
25.8781 | -105.9672
25.8818 | -105.9638
25.8855 | -105.9625
25.8892 | -105.9613
25.8928 | -105.9601
25.8965 | -105.9592
25.9002 | -105.9575
25.9039 | -105.9574
25.9075 | -105.9573
25.9112 | -105.9552
25.9149 | -105.9537
25.9185 | -105.9531
25.9222 | -105.9528
25.9259 | -105.9524
25.9296 | -105.9507
25.9332 | -105.9488
25.9369 | -105.9465
25.9406 | -105.944
25.9443 | -105.9417
25.9479 | -105.9414
25.9516 | -105.9416
25.9553 | -105.9411
25.959 | -105.9413
25.9626 | -105.9408
25.9663 | -105.9384
25.97 | -105.9355
25.9736 | -105.9337
25.9773 | -105.9335
25.981 | -105.9354
25.9847 | -105.9344
25.9883 | -105.9321
25.992 | -105.931
25.9957 | -105.9289
25.9994 | -105.9275
26.003 | -105.9274
26.0067 | -105.9262
26.0104 | -105.9247
26.0141 | -105.9228
26.0177 | -105.9214
26.0214 | -105.9201
26.0251 | -105.9178
26.0287 | -105.9164
26.0324 | -105.9153
26.0361 | -105.9119
26.0398 | -105.9099
26.0434 | -105.9066
26.0471 | -105.9024
26.0508 | -105.8988
26.0545 | -105.8964
26.0581 | -105.8936
26.0618 | -105.8906
26.0655 | -105.8886
26.0691 | -105.8864
26.0728 | -105.8829
26.0765 | -105.878
26.0802 | -105.8736
26.0838 | -105.8695
26.0875 | -105.8637
26.0912 | -105.857
26.0949 | -105.8514
26.0985 | -105.8465
26.1022 | -105.8413
26.1059 | -105.8363
26.1096 | -105.8325
26.1132 | -105.8275
26.1169 | -105.8237
26.1206 | -105.8188
26.1242 | -105.8136
26.1279 | -105.8098
26.1316 | -105.8052
26.1353 | -105.8012
26.1389 | -105.7961
26.1426 | -105.7921
26.1463 | -105.7888
26.15 | -105.7844
26.1536 | -105.7802
26.1573 | -105.7772
26.161 | -105.7738
26.1647 | -105.771
26.1683 | -105.7675
26.172 | -105.765
26.1757 | -105.7637
26.1793 | -105.7619
26.183 | -105.7598
26.1867 | -105.7584
26.1904 | -105.7566
26.194 | -105.7545
26.1977 | -105.7521
26.2014 | -105.7515
26.2051 | -105.75
26.2087 | -105.7501
26.2124 | -105.751
26.2161 | -105.7522
26.2198 | -105.7541
26.2234 | -105.756
26.2271 | -105.7602
26.2308 | -105.7639
26.2344 | -105.7666
26.2381 | -105.7678
26.2418 | -105.7658
26.2455 | -105.7624
26.2491 | -105.7618
26.2528 | -105.7657
26.2565 | -105.7687
26.2602 | -105.7729
26.2638 | -105.7783
26.2675 | -105.7835
26.2712 | -105.7878
26.2749 | -105.7945
26.2785 | -105.8031
26.2822 | -105.812
26.2859 | -105.8202
26.2895 | -105.828
26.2932 | -105.8356
26.2969 | -105.8425
26.3006 | -105.85
26.3042 | -105.8584
26.3079 | -105.8654
26.3116 | -105.8733
26.3153 | -105.8802
26.3189 | -105.8877
26.3226 | -105.8937
26.3263 | -105.8997
26.33 | -105.9078
26.3336 | -105.9155
26.3373 | -105.9227
26.341 | -105.9304
26.3446 | -105.9364
26.3483 | -105.9402
26.352 | -105.9445
26.3557 | -105.9509
26.3593 | -105.9573
26.363 | -105.9623
26.3667 | -105.9683
26.3704 | -105.9735
26.374 | -105.9772
26.3777 | -105.9819
26.3814 | -105.988
26.385 | -105.9934
26.3887 | -105.9976
26.3924 | -106.0017
26.3961 | -106.0076
26.3997 | -106.0134
26.4034 | -106.0185
26.4071 | -106.0256
26.4108 | -106.0313
26.4144 | -106.0386
26.4181 | -106.0468
26.4218 | -106.0555
26.4255 | -106.0644
26.4291 | -106.0734
26.4328 | -106.0821
26.4365 | -106.091
26.4401 | -106.1002
26.4438 | -106.1102
26.4475 | -106.1209
26.4512 | -106.132
26.4548 | -106.1431
26.4585 | -106.1566
26.4622 | -106.1728
26.4659 | -106.1872
26.4695 | -106.2008
26.4732 | -106.2149
26.4769 | -106.2294
26.4806 | -106.2441
26.4842 | -106.259
26.4879 | -106.2735
26.4916 | -106.2899
26.4952 | -106.3053
26.4989 | -106.3197
26.5026 | -106.333
26.5063 | -106.3467
26.5099 | -106.3574
26.5136 | -106.3693
26.5173 | -106.3797
26.521 | -106.3877
26.5246 | -106.3971
26.5283 | -106.4094
26.532 | -106.4205
26.5357 | -106.4299
26.5393 | -106.439
26.543 | -106.448
26.5467 | -106.4576
26.5503 | -106.4656
26.554 | -106.4727
26.5577 | -106.4803
26.5614 | -106.4868
26.565 | -106.4921
26.5687 | -106.4981
26.5724 | -106.5061
26.5761 | -106.5124
26.5797 | -106.518
26.5834 | -106.5219
26.5871 | -106.5244
26.5908 | -106.5265
26.5944 | -106.5277
26.5981 | -106.529
26.6018 | -106.5279
26.6054 | -106.5272
26.6091 | -106.5261
26.6128 | -106.5255
26.6165 | -106.5305
26.6201 | -106.538
26.6238 | -106.54
26.6275 | -106.5377
26.6312 | -106.5357
26.6348 | -106.5323
26.6385 | -106.5296
26.6422 | -106.5269
26.6459 | -106.5245
26.6495 | -106.522
26.6532 | -106.5181
26.6569 | -106.5134
26.6605 | -106.509
26.6642 | -106.5062
26.6679 | -106.5052
26.6716 | -106.5046
26.6752 | -106.5027
26.6789 | -106.5009
26.6826 | -106.5012
26.6863 | -106.5019
26.6899 | -106.501
26.6936 | -106.5005
26.6973 | -106.503
26.7009 | -106.5061
26.7046 | -106.5051
26.7083 | -106.5048
26.712 | -106.504
26.7156 | -106.5029
26.7193 | -106.5043
26.723 | -106.5069
26.7267 | -106.5087
26.7303 | -106.5104
26.734 | -106.5115
26.7377 | -106.5138
26.7414 | -106.5172
26.745 | -106.5189
26.7487 | -106.5206
26.7524 | -106.5222
26.756 | -106.5235
26.7597 | -106.5264
26.7634 | -106.5297
26.7671 | -106.5327
26.7707 | -106.5344
26.7744 | -106.5352
26.7781 | -106.5374
26.7818 | -106.5392
26.7854 | -106.5405
26.7891 | -106.5413
26.7928 | -106.5412
26.7965 | -106.5412
26.8001 | -106.5415
26.8038 | -106.5393
26.8075 | -106.5357
26.8111 | -106.5346
26.8148 | -106.5322
26.8185 | -106.5304
26.8222 | -106.5286
26.8258 | -106.5242
26.8295 | -106.5183
26.8332 | -106.5128
26.8369 | -106.506
26.8405 | -106.5004
26.8442 | -106.4935
26.8479 | -106.4868
26.8516 | -106.4817
26.8552 | -106.4765
26.8589 | -106.4694
26.8626 | -106.4627
26.8662 | -106.4566
26.8699 | -106.451
26.8736 | -106.4459
26.8773 | -106.4414
26.8809 | -106.4372
26.8846 | -106.434
26.8883 | -106.4316
26.892 | -106.4285
26.8956 | -106.4261
26.8993 | -106.4242
26.903 | -106.4191
26.9067 | -106.4134
26.9103 | -106.4103
26.914 | -106.4075
26.9177 | -106.4031
26.9213 | -106.3995
26.925 | -106.397
26.9287 | -106.3946
26.9324 | -106.3944
26.936 | -106.3934
26.9397 | -106.3922
26.9434 | -106.3918
26.9471 | -106.3911
26.9507 | -106.3917
26.9544 | -106.3923
26.9581 | -106.3931
26.9618 | -106.3954
26.9654 | -106.3986
26.9691 | -106.4014
26.9728 | -106.4033
26.9764 | -106.4064
26.9801 | -106.4101
26.9838 | -106.4143
26.9875 | -106.4208
26.9911 | -106.4275
26.9948 | -106.432
26.9985 | -106.4369
27.0022 | -106.4415
27.0058 | -106.4428
27.0095 | -106.4461
27.0132 | -106.4501
27.0169 | -106.4535
27.0205 | -106.4567
27.0242 | -106.4585
27.0279 | -106.4619
27.0315 | -106.4655
27.0352 | -106.4701
27.0389 | -106.4742
27.0426 | -106.4791
27.0462 | -106.4835
27.0499 | -106.488
27.0536 | -106.4911
27.0573 | -106.4946
27.0609 | -106.4985
27.0646 | -106.5024
27.0683 | -106.5065
27.072 | -106.5113
27.0756 | -106.5156
27.0793 | -106.5182
27.083 | -106.5218
27.0866 | -106.5267
27.0903 | -106.532
27.094 | -106.5392
27.0977 | -106.5459
27.1013 | -106.549
27.105 | -106.5535
27.1087 | -106.5575
27.1124 | -106.5616
27.116 | -106.5664
27.1197 | -106.5702
27.1234 | -106.5725
27.127 | -106.5769
27.1307 | -106.5811
27.1344 | -106.5856
27.1381 | -106.5893
27.1417 | -106.5931
27.1454 | -106.5973
27.1491 | -106.6016
27.1528 | -106.6055
27.1564 | -106.6093
27.1601 | -106.6117
27.1638 | -106.6157
27.1675 | -106.6213
27.1711 | -106.6249
27.1748 | -106.6282
27.1785 | -106.6304
27.1821 | -106.6332
27.1858 | -106.637
27.1895 | -106.6424
27.1932 | -106.6453
27.1968 | -106.6484
27.2005 | -106.6509
27.2042 | -106.6517
27.2079 | -106.6551
27.2115 | -106.6591
27.2152 | -106.6613
27.2189 | -106.6635
27.2226 | -106.6652
27.2262 | -106.6669
27.2299 | -106.6687
27.2336 | -106.667
27.2372 | -106.6671
27.2409 | -106.6693
27.2446 | -106.6708
27.2483 | -106.6694
27.2519 | -106.6689
27.2556 | -106.6659
27.2593 | -106.6626
27.263 | -106.6606
27.2666 | -106.6587
27.2703 | -106.6567
27.274 | -106.6551
27.2777 | -106.6536
27.2813 | -106.6521
27.285 | -106.6495
27.2887 | -106.6447
27.2923 | -106.6394
27.296 | -106.635
27.2997 | -106.6311
27.3034 | -106.6268
27.307 | -106.6215
27.3107 | -106.6174
27.3144 | -106.6123
27.3181 | -106.6063
27.3217 | -106.5992
27.3254 | -106.5946
27.3291 | -106.588
27.3328 | -106.5798
27.3364 | -106.572
27.3401 | -106.5671
27.3438 | -106.5625
27.3474 | -106.5569
27.3511 | -106.55
27.3548 | -106.544
27.3585 | -106.5387
27.3621 | -106.5332
27.3658 | -106.5272
27.3695 | -106.5209
27.3732 | -106.514
27.3768 | -106.5083
27.3805 | -106.5033
27.3842 | -106.499
27.3879 | -106.4965
27.3915 | -106.4916
27.3952 | -106.4873
27.3989 | -106.4835
27.4025 | -106.4801
27.4062 | -106.478
27.4099 | -106.4744
27.4136 | -106.4707
27.4172 | -106.4661
27.4209 | -106.4617
27.4246 | -106.4599
27.4283 | -106.4582
27.4319 | -106.4556
27.4356 | -106.4522
27.4393 | -106.4509
27.4429 | -106.4495
27.4466 | -106.4476
27.4503 | -106.4469
27.454 | -106.4464
27.4576 | -106.4474
27.4613 | -106.4485
27.465 | -106.4497
27.4687 | -106.4502
27.4723 | -106.4512
27.476 | -106.4518
27.4797 | -106.4522
27.4834 | -106.4517
27.487 | -106.4537
27.4907 | -106.4551
27.4944 | -106.4563
27.498 | -106.4579
27.5017 | -106.4586
27.5054 | -106.4606
27.5091 | -106.4623
27.5127 | -106.4633
27.5164 | -106.4658
27.5201 | -106.469
27.5238 | -106.473
27.5274 | -106.4752
27.5311 | -106.4765
27.5348 | -106.478
27.5385 | -106.4793
27.5421 | -106.4803
27.5458 | -106.4823
27.5495 | -106.4834
27.5531 | -106.4809
27.5568 | -106.4805
27.5605 | -106.4821
27.5642 | -106.4838
27.5678 | -106.4846
27.5715 | -106.4847
27.5752 | -106.4854
27.5789 | -106.4838
27.5825 | -106.483
27.5862 | -106.4818
27.5899 | -106.4786
27.5936 | -106.4767
27.5972 | -106.4747
27.6009 | -106.4708
27.6046 | -106.4666
27.6082 | -106.4612
27.6119 | -106.4577
27.6156 | -106.4557
27.6193 | -106.452
27.6229 | -106.4482
27.6266 | -106.4426
27.6303 | -106.4379
27.634 | -106.4338
27.6376 | -106.4292
27.6413 | -106.4226
27.645 | -106.4168
27.6487 | -106.4105
27.6523 | -106.4035
27.656 | -106.3966
27.6597 | -106.3905
27.6633 | -106.3837
27.667 | -106.3761
27.6707 | -106.3687
27.6744 | -106.3633
27.678 | -106.357
27.6817 | -106.3493
27.6854 | -106.3411
27.6891 | -106.3318
27.6927 | -106.3235
27.6964 | -106.3163
27.7001 | -106.3086
27.7038 | -106.2995
27.7074 | -106.2935
27.7111 | -106.2883
27.7148 | -106.28
27.7184 | -106.2727
27.7221 | -106.2653
27.7258 | -106.2574
27.7295 | -106.2477
27.7331 | -106.2385
27.7368 | -106.2284
27.7405 | -106.2186
27.7442 | -106.2113
27.7478 | -106.202
27.7515 | -106.1934
27.7552 | -106.1865
27.7588 | -106.1786
27.7625 | -106.1728
27.7662 | -106.166
27.7699 | -106.1578
27.7735 | -106.1507
27.7772 | -106.1427
27.7809 | -106.1355
27.7846 | -106.1285
27.7882 | -106.1216
27.7919 | -106.1154
27.7956 | -106.1077
27.7993 | -106.1006
27.8029 | -106.0956
27.8066 | -106.0894
27.8103 | -106.081
27.8139 | -106.0733
27.8176 | -106.0667
27.8213 | -106.0617
27.825 | -106.057
27.8286 | -106.0523
27.8323 | -106.046
27.836 | -106.0415
27.8397 | -106.0365
27.8433 | -106.0317
27.847 | -106.0264
27.8507 | -106.0219
27.8544 | -106.0169
27.858 | -106.008
27.8617 | -106.002
27.8654 | -105.9961
27.869 | -105.991
27.8727 | -105.9877
27.8764 | -105.9834
27.8801 | -105.9787
27.8837 | -105.9752
27.8874 | -105.9704
27.8911 | -105.9658
27.8948 | -105.9617
27.8984 | -105.9594
27.9021 | -105.9567
27.9058 | -105.9528
27.9095 | -105.9471
27.9131 | -105.9417
27.9168 | -105.9377
27.9205 | -105.934
27.9241 | -105.9238
27.9278 | -105.901
27.9315 | -105.8781
27.9352 | -105.865
27.9388 | -105.8537
27.9425 | -105.8432
27.9462 | -105.8324
27.9499 | -105.8199
27.9535 | -105.8063
27.9572 | -105.7927
27.9609 | -105.7803
27.9646 | -105.7675
27.9682 | -105.7526
27.9719 | -105.7374
27.9756 | -105.7211
27.9792 | -105.7028
27.9829 | -105.6859
27.9866 | -105.6678
27.9903 | -105.6487
27.9939 | -105.6294
27.9976 | -105.6081
28.0013 | -105.5861