## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 277 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 283
Hood | 02-92B/P | 138
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1008.2 | hPa |  | PRC PRO 01
Humidity | 58.8 | %/Hr
Internal temperature of calibration bench | 24.22 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1011 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.4 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.02 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.98 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3112 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 23/08/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 28/08/2018
Technical Manager | Anthony HUE | Approval |  | 30/08/2018
CEO | Robert GLEMOT | Approval |  | 30/08/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4317 | 57.4349
0.012209 | -36.0057 | 38.2473
0.018313 | -34.9426 | 28.0855
0.024417 | -34.4862 | 21.8497
0.030521 | -34.3752 | 18.1057
0.036626 | -34.2335 | 15.0504
0.04273 | -34.1397 | 12.6835
0.048834 | -34.0948 | 11.1731
0.054938 | -34.0305 | 10.0328
0.061043 | -33.9711 | 8.945
0.067147 | -33.9622 | 8.1191
0.073251 | -33.9646 | 7.3489
0.079355 | -33.9838 | 6.6725
0.08546 | -33.9975 | 6.0693
0.091564 | -33.9924 | 5.6371
0.097668 | -33.9751 | 5.4039
0.10377 | -33.9557 | 5.1565
0.10988 | -33.941 | 4.8821
0.11598 | -33.9341 | 4.5915
0.12209 | -33.9374 | 4.3092
0.12819 | -33.9476 | 4.0229
0.13429 | -33.9576 | 3.7322
0.1404 | -33.9604 | 3.4567
0.1465 | -33.9624 | 3.1984
0.15261 | -33.9602 | 2.9728
0.15871 | -33.9543 | 2.7997
0.16482 | -33.9464 | 2.6841
0.17092 | -33.9443 | 2.6411
0.17702 | -33.9423 | 2.567
0.18313 | -33.9414 | 2.4918
0.18923 | -33.9408 | 2.387
0.19534 | -33.9392 | 2.2639
0.20144 | -33.9376 | 2.1215
0.20754 | -33.9349 | 1.9955
0.21365 | -33.934 | 1.8785
0.21975 | -33.9337 | 1.7787
0.22586 | -33.9348 | 1.6918
0.23196 | -33.9373 | 1.6226
0.23807 | -33.9413 | 1.5643
0.24417 | -33.9455 | 1.5089
0.25027 | -33.9498 | 1.4481
0.25638 | -33.9525 | 1.3873
0.26248 | -33.9547 | 1.3196
0.26859 | -33.9559 | 1.2525
0.27469 | -33.9567 | 1.1783
0.2808 | -33.9578 | 1.1023
0.2869 | -33.9587 | 1.03
0.293 | -33.9597 | 0.96177
0.29911 | -33.9607 | 0.901
0.30521 | -33.9617 | 0.84253
0.31132 | -33.9629 | 0.80614
0.31742 | -33.9646 | 0.76451
0.32353 | -33.9655 | 0.73162
0.32963 | -33.9672 | 0.70474
0.33573 | -33.9693 | 0.68618
0.34184 | -33.9716 | 0.66574
0.34794 | -33.9747 | 0.64078
0.35405 | -33.9777 | 0.61666
0.36015 | -33.981 | 0.59044
0.36626 | -33.9837 | 0.56133
0.37236 | -33.9863 | 0.52245
0.37846 | -33.9888 | 0.47463
0.38457 | -33.9895 | 0.41314
0.39067 | -33.9905 | 0.36358
0.39678 | -33.991 | 0.30786
0.40288 | -33.9913 | 0.25036
0.40899 | -33.991 | 0.19491
0.41509 | -33.9901 | 0.14187
0.42119 | -33.9901 | 0.098195
0.4273 | -33.99 | 0.049986
0.4334 | -33.9899 | 0.002705
0.43951 | -33.9904 | -0.040888
0.44561 | -33.9915 | -0.081144
0.45172 | -33.9936 | -0.11184
0.45782 | -33.9948 | -0.14841
0.46392 | -33.9966 | -0.18445
0.47003 | -33.9981 | -0.21562
0.47613 | -33.9994 | -0.24953
0.48224 | -34.0005 | -0.28376
0.48834 | -34 | -0.31884
0.49445 | -34.0002 | -0.35355
0.50055 | -34.0001 | -0.38538
0.50665 | -33.9998 | -0.41941
0.51276 | -33.999 | -0.45396
0.51886 | -33.9983 | -0.4886
0.52497 | -33.9972 | -0.5133
0.53107 | -33.9966 | -0.5436
0.53717 | -33.996 | -0.57212
0.54328 | -33.9952 | -0.59545
0.54938 | -33.9951 | -0.61843
0.55549 | -33.9951 | -0.63992
0.56159 | -33.9961 | -0.65201
0.5677 | -33.9963 | -0.66711
0.5738 | -33.9972 | -0.68426
0.5799 | -33.9977 | -0.69973
0.58601 | -33.9986 | -0.70993
0.59211 | -34 | -0.71634
0.59822 | -34.0011 | -0.72604
0.60432 | -34.0023 | -0.7391
0.61043 | -34.0033 | -0.74732
0.61653 | -34.0044 | -0.75835
0.62263 | -34.0057 | -0.76734
0.62874 | -34.0064 | -0.77787
0.63484 | -34.0071 | -0.78703
0.64095 | -34.0076 | -0.79731
0.64705 | -34.008 | -0.80716
0.65316 | -34.0083 | -0.81938
0.65926 | -34.0086 | -0.8348
0.66536 | -34.0079 | -0.85163
0.67147 | -34.0078 | -0.86458
0.67757 | -34.0078 | -0.87883
0.68368 | -34.0074 | -0.89794
0.68978 | -34.007 | -0.91473
0.69589 | -34.0064 | -0.934
0.70199 | -34.0058 | -0.95722
0.70809 | -34.0053 | -0.97627
0.7142 | -34.005 | -0.99729
0.7203 | -34.0043 | -1.0168
0.72641 | -34.0039 | -1.04
0.73251 | -34.0036 | -1.0606
0.73862 | -34.0041 | -1.0837
0.74472 | -34.004 | -1.1043
0.75082 | -34.0039 | -1.1275
0.75693 | -34.0045 | -1.1505
0.76303 | -34.0052 | -1.1739
0.76914 | -34.0072 | -1.1962
0.77524 | -34.0084 | -1.217
0.78135 | -34.0099 | -1.2387
0.78745 | -34.0114 | -1.2604
0.79355 | -34.0128 | -1.282
0.79966 | -34.0147 | -1.3031
0.80576 | -34.0171 | -1.325
0.81187 | -34.0189 | -1.3485
0.81797 | -34.021 | -1.3717
0.82408 | -34.0229 | -1.3943
0.83018 | -34.0248 | -1.4166
0.83628 | -34.0269 | -1.4371
0.84239 | -34.0283 | -1.4591
0.84849 | -34.0298 | -1.4826
0.8546 | -34.0314 | -1.504
0.8607 | -34.0328 | -1.5265
0.86681 | -34.034 | -1.5467
0.87291 | -34.0352 | -1.5667
0.87901 | -34.0358 | -1.5853
0.88512 | -34.0367 | -1.6043
0.89122 | -34.0375 | -1.6233
0.89733 | -34.038 | -1.6397
0.90343 | -34.0385 | -1.6577
0.90953 | -34.0387 | -1.6716
0.91564 | -34.0388 | -1.6878
0.92174 | -34.0391 | -1.703
0.92785 | -34.0392 | -1.7176
0.93395 | -34.0393 | -1.7313
0.94006 | -34.0395 | -1.7441
0.94616 | -34.0392 | -1.7558
0.95226 | -34.0394 | -1.7684
0.95837 | -34.0396 | -1.7823
0.96447 | -34.0398 | -1.7948
0.97058 | -34.0397 | -1.808
0.97668 | -34.0395 | -1.821
0.98279 | -34.0393 | -1.8336
0.98889 | -34.0392 | -1.844
0.99499 | -34.0393 | -1.8574
1.0011 | -34.0392 | -1.8678
1.0072 | -34.039 | -1.8793
1.0133 | -34.0387 | -1.8899
1.0194 | -34.0384 | -1.9031
1.0255 | -34.0382 | -1.9169
1.0316 | -34.0377 | -1.9289
1.0377 | -34.0374 | -1.9429
1.0438 | -34.037 | -1.9556
1.0499 | -34.0363 | -1.9711
1.056 | -34.0357 | -1.9856
1.0621 | -34.0352 | -2.0001
1.0682 | -34.0348 | -2.0127
1.0743 | -34.0343 | -2.0285
1.0805 | -34.0339 | -2.043
1.0866 | -34.0331 | -2.0586
1.0927 | -34.0324 | -2.0734
1.0988 | -34.0318 | -2.0883
1.1049 | -34.0311 | -2.1027
1.111 | -34.0304 | -2.1183
1.1171 | -34.0298 | -2.1333
1.1232 | -34.0294 | -2.1487
1.1293 | -34.029 | -2.163
1.1354 | -34.0289 | -2.1782
1.1415 | -34.0286 | -2.1941
1.1476 | -34.028 | -2.2099
1.1537 | -34.0276 | -2.2249
1.1598 | -34.0277 | -2.2399
1.1659 | -34.0276 | -2.2537
1.172 | -34.0274 | -2.2678
1.1781 | -34.0272 | -2.2818
1.1842 | -34.0271 | -2.296
1.1903 | -34.0275 | -2.3069
1.1964 | -34.0274 | -2.3208
1.2025 | -34.0276 | -2.3349
1.2086 | -34.028 | -2.3475
1.2147 | -34.0281 | -2.3602
1.2209 | -34.0285 | -2.3714
1.227 | -34.0293 | -2.3821
1.2331 | -34.0297 | -2.3941
1.2392 | -34.0301 | -2.4058
1.2453 | -34.0306 | -2.4169
1.2514 | -34.031 | -2.4267
1.2575 | -34.0316 | -2.4376
1.2636 | -34.0325 | -2.4463
1.2697 | -34.0331 | -2.4554
1.2758 | -34.0339 | -2.4665
1.2819 | -34.0346 | -2.4762
1.288 | -34.0354 | -2.4869
1.2941 | -34.0362 | -2.4953
1.3002 | -34.037 | -2.5032
1.3063 | -34.0376 | -2.512
1.3124 | -34.0384 | -2.5211
1.3185 | -34.0392 | -2.5297
1.3246 | -34.0399 | -2.5384
1.3307 | -34.0408 | -2.5468
1.3368 | -34.0417 | -2.5553
1.3429 | -34.0424 | -2.564
1.349 | -34.043 | -2.5733
1.3551 | -34.0438 | -2.5824
1.3613 | -34.0445 | -2.5904
1.3674 | -34.0452 | -2.6003
1.3735 | -34.0457 | -2.6101
1.3796 | -34.0462 | -2.619
1.3857 | -34.0466 | -2.6284
1.3918 | -34.0472 | -2.6385
1.3979 | -34.0479 | -2.6477
1.404 | -34.0485 | -2.6593
1.4101 | -34.049 | -2.6699
1.4162 | -34.0495 | -2.6808
1.4223 | -34.0501 | -2.6922
1.4284 | -34.0506 | -2.7023
1.4345 | -34.051 | -2.7141
1.4406 | -34.0514 | -2.7281
1.4467 | -34.0519 | -2.74
1.4528 | -34.0525 | -2.7517
1.4589 | -34.0531 | -2.764
1.465 | -34.0534 | -2.7751
1.4711 | -34.054 | -2.7867
1.4772 | -34.0547 | -2.8006
1.4833 | -34.0552 | -2.8141
1.4894 | -34.0559 | -2.8267
1.4955 | -34.0565 | -2.8389
1.5016 | -34.0572 | -2.8518
1.5078 | -34.0577 | -2.8667
1.5139 | -34.0582 | -2.8802
1.52 | -34.0588 | -2.8935
1.5261 | -34.0593 | -2.9068
1.5322 | -34.0597 | -2.9213
1.5383 | -34.0603 | -2.9344
1.5444 | -34.061 | -2.9479
1.5505 | -34.0616 | -2.962
1.5566 | -34.0623 | -2.9756
1.5627 | -34.0629 | -2.99
1.5688 | -34.0636 | -3.0042
1.5749 | -34.0641 | -3.0176
1.581 | -34.0646 | -3.0317
1.5871 | -34.0651 | -3.0462
1.5932 | -34.0657 | -3.0598
1.5993 | -34.0662 | -3.0747
1.6054 | -34.0668 | -3.0886
1.6115 | -34.0674 | -3.1035
1.6176 | -34.0678 | -3.1166
1.6237 | -34.0682 | -3.131
1.6298 | -34.0686 | -3.1441
1.6359 | -34.0692 | -3.1576
1.642 | -34.0698 | -3.1701
1.6482 | -34.0701 | -3.1838
1.6543 | -34.0706 | -3.1981
1.6604 | -34.0711 | -3.2108
1.6665 | -34.0714 | -3.2238
1.6726 | -34.0717 | -3.2368
1.6787 | -34.0722 | -3.2502
1.6848 | -34.0723 | -3.2634
1.6909 | -34.0727 | -3.2769
1.697 | -34.073 | -3.2898
1.7031 | -34.0732 | -3.3031
1.7092 | -34.0733 | -3.3159
1.7153 | -34.0736 | -3.3283
1.7214 | -34.0735 | -3.3409
1.7275 | -34.0735 | -3.354
1.7336 | -34.0736 | -3.3675
1.7397 | -34.0738 | -3.3802
1.7458 | -34.0739 | -3.393
1.7519 | -34.074 | -3.4058
1.758 | -34.0741 | -3.4182
1.7641 | -34.0744 | -3.4301
1.7702 | -34.0744 | -3.4436
1.7763 | -34.0744 | -3.4555
1.7824 | -34.0745 | -3.4685
1.7885 | -34.0744 | -3.4811
1.7947 | -34.0746 | -3.4935
1.8008 | -34.0747 | -3.5052
1.8069 | -34.0747 | -3.5171
1.813 | -34.0749 | -3.5296
1.8191 | -34.0748 | -3.5427
1.8252 | -34.0749 | -3.5559
1.8313 | -34.0751 | -3.5685
1.8374 | -34.0753 | -3.5806
1.8435 | -34.0753 | -3.5929
1.8496 | -34.0754 | -3.6062
1.8557 | -34.0756 | -3.6191
1.8618 | -34.0759 | -3.6315
1.8679 | -34.076 | -3.6439
1.874 | -34.0761 | -3.6563
1.8801 | -34.0763 | -3.6689
1.8862 | -34.0765 | -3.6813
1.8923 | -34.0766 | -3.6938
1.8984 | -34.0768 | -3.7062
1.9045 | -34.0768 | -3.7179
1.9106 | -34.0769 | -3.7303
1.9167 | -34.0769 | -3.7427
1.9228 | -34.0771 | -3.7549
1.9289 | -34.0772 | -3.7667
1.9351 | -34.0774 | -3.7775
1.9412 | -34.0777 | -3.789
1.9473 | -34.078 | -3.8013
1.9534 | -34.0782 | -3.8129
1.9595 | -34.0785 | -3.8249
1.9656 | -34.0788 | -3.8373
1.9717 | -34.0791 | -3.8491
1.9778 | -34.0794 | -3.8607
1.9839 | -34.0797 | -3.8723
1.99 | -34.08 | -3.8834
1.9961 | -34.0801 | -3.8957
2.0022 | -34.0804 | -3.9072
2.0083 | -34.0805 | -3.9185
2.0144 | -34.0808 | -3.9301
2.0205 | -34.0812 | -3.942
2.0266 | -34.0816 | -3.9535
2.0327 | -34.0819 | -3.9653
2.0388 | -34.0824 | -3.9769
2.0449 | -34.0827 | -3.9885
2.051 | -34.083 | -4.0001
2.0571 | -34.0834 | -4.0115
2.0632 | -34.0837 | -4.0231
2.0693 | -34.0839 | -4.0353
2.0754 | -34.0843 | -4.0469
2.0816 | -34.0847 | -4.0591
2.0877 | -34.085 | -4.0708
2.0938 | -34.0854 | -4.083
2.0999 | -34.0857 | -4.095
2.106 | -34.086 | -4.1068
2.1121 | -34.0863 | -4.1189
2.1182 | -34.0866 | -4.1297
2.1243 | -34.087 | -4.1405
2.1304 | -34.0873 | -4.1518
2.1365 | -34.0877 | -4.1638
2.1426 | -34.088 | -4.1761
2.1487 | -34.0884 | -4.1874
2.1548 | -34.0889 | -4.1993
2.1609 | -34.0894 | -4.2107
2.167 | -34.0899 | -4.2221
2.1731 | -34.0903 | -4.2342
2.1792 | -34.0906 | -4.2454
2.1853 | -34.091 | -4.2574
2.1914 | -34.0914 | -4.2689
2.1975 | -34.0919 | -4.2806
2.2036 | -34.0922 | -4.2927
2.2097 | -34.0925 | -4.3043
2.2158 | -34.093 | -4.315
2.222 | -34.0934 | -4.3255
2.2281 | -34.0937 | -4.337
2.2342 | -34.0941 | -4.3485
2.2403 | -34.0944 | -4.3593
2.2464 | -34.0947 | -4.3702
2.2525 | -34.095 | -4.3817
2.2586 | -34.0955 | -4.3924
2.2647 | -34.0959 | -4.4029
2.2708 | -34.0962 | -4.4144
2.2769 | -34.0967 | -4.425
2.283 | -34.0969 | -4.4365
2.2891 | -34.0972 | -4.4475
2.2952 | -34.0974 | -4.4582
2.3013 | -34.0977 | -4.4692
2.3074 | -34.098 | -4.4797
2.3135 | -34.0983 | -4.4914
2.3196 | -34.0985 | -4.5016
2.3257 | -34.0988 | -4.5124
2.3318 | -34.0989 | -4.5238
2.3379 | -34.0992 | -4.5346
2.344 | -34.0994 | -4.5461
2.3501 | -34.0996 | -4.5569
2.3562 | -34.0998 | -4.5674
2.3623 | -34.0999 | -4.5779
2.3685 | -34.1001 | -4.5886
2.3746 | -34.1002 | -4.5999
2.3807 | -34.1005 | -4.6106
2.3868 | -34.1006 | -4.6215
2.3929 | -34.1008 | -4.6323
2.399 | -34.1009 | -4.6428
2.4051 | -34.1012 | -4.6534
2.4112 | -34.1014 | -4.6642
2.4173 | -34.1016 | -4.675
2.4234 | -34.1016 | -4.686
2.4295 | -34.1018 | -4.697
2.4356 | -34.102 | -4.708
2.4417 | -34.1021 | -4.7192
2.4478 | -34.1022 | -4.7303
2.4539 | -34.1024 | -4.7407
2.46 | -34.1025 | -4.7524
2.4661 | -34.1026 | -4.7636
2.4722 | -34.1028 | -4.774
2.4783 | -34.1028 | -4.7846
2.4844 | -34.1029 | -4.7959
2.4905 | -34.103 | -4.8069
2.4966 | -34.1032 | -4.8179
2.5027 | -34.1033 | -4.8285
2.5089 | -34.1034 | -4.8396
2.515 | -34.1034 | -4.8505
2.5211 | -34.1035 | -4.8614
2.5272 | -34.1036 | -4.8725
2.5333 | -34.1037 | -4.8837
2.5394 | -34.1039 | -4.8946
2.5455 | -34.104 | -4.9061
2.5516 | -34.1041 | -4.9167
2.5577 | -34.1043 | -4.9283
2.5638 | -34.1044 | -4.9402
2.5699 | -34.1045 | -4.9517
2.576 | -34.1045 | -4.9635
2.5821 | -34.1046 | -4.975
2.5882 | -34.1047 | -4.9864
2.5943 | -34.1049 | -4.9973
2.6004 | -34.1051 | -5.0086
2.6065 | -34.1052 | -5.0191
2.6126 | -34.1054 | -5.0296
2.6187 | -34.1055 | -5.0402
2.6248 | -34.1056 | -5.0512
2.6309 | -34.1058 | -5.0629
2.637 | -34.1061 | -5.0748
2.6431 | -34.1063 | -5.0854
2.6492 | -34.1064 | -5.0971
2.6554 | -34.1067 | -5.1081
2.6615 | -34.1068 | -5.1193
2.6676 | -34.1071 | -5.1309
2.6737 | -34.1074 | -5.1423
2.6798 | -34.1077 | -5.1536
2.6859 | -34.1079 | -5.1651
2.692 | -34.1081 | -5.1761
2.6981 | -34.1084 | -5.1874
2.7042 | -34.1086 | -5.1987
2.7103 | -34.1089 | -5.2095
2.7164 | -34.1092 | -5.2207
2.7225 | -34.1095 | -5.2323
2.7286 | -34.1098 | -5.2435
2.7347 | -34.11 | -5.2542
2.7408 | -34.1104 | -5.2644
2.7469 | -34.1106 | -5.2758
2.753 | -34.111 | -5.2867
2.7591 | -34.1113 | -5.2982
2.7652 | -34.1116 | -5.31
2.7713 | -34.1118 | -5.3209
2.7774 | -34.1122 | -5.332
2.7835 | -34.1124 | -5.3431
2.7896 | -34.1126 | -5.3538
2.7958 | -34.113 | -5.3647
2.8019 | -34.1133 | -5.3753
2.808 | -34.1137 | -5.3863
2.8141 | -34.1141 | -5.3977
2.8202 | -34.1144 | -5.4088
2.8263 | -34.1148 | -5.4198
2.8324 | -34.1152 | -5.4314
2.8385 | -34.1157 | -5.4424
2.8446 | -34.1161 | -5.4533
2.8507 | -34.1166 | -5.4644
2.8568 | -34.117 | -5.475
2.8629 | -34.1173 | -5.4858
2.869 | -34.1177 | -5.4971
2.8751 | -34.1181 | -5.508
2.8812 | -34.1186 | -5.5195
2.8873 | -34.119 | -5.5305
2.8934 | -34.1194 | -5.5417
2.8995 | -34.1198 | -5.553
2.9056 | -34.1201 | -5.5644
2.9117 | -34.1205 | -5.5756
2.9178 | -34.121 | -5.5866
2.9239 | -34.1214 | -5.5982
2.93 | -34.1219 | -5.6098
2.9361 | -34.1223 | -5.6208
2.9423 | -34.1227 | -5.6317
2.9484 | -34.1232 | -5.643
2.9545 | -34.1236 | -5.6542
2.9606 | -34.1241 | -5.6657
2.9667 | -34.1245 | -5.6771
2.9728 | -34.1249 | -5.6886
2.9789 | -34.1254 | -5.6999
2.985 | -34.1259 | -5.7114
2.9911 | -34.1263 | -5.7229
2.9972 | -34.1269 | -5.7342
3.0033 | -34.1273 | -5.7459
3.0094 | -34.1277 | -5.7571
3.0155 | -34.1282 | -5.7683
3.0216 | -34.1286 | -5.7798
3.0277 | -34.129 | -5.7916
3.0338 | -34.1295 | -5.8032
3.0399 | -34.1299 | -5.8148
3.046 | -34.1303 | -5.8266
3.0521 | -34.1307 | -5.8378
3.0582 | -34.131 | -5.8502
3.0643 | -34.1315 | -5.8617
3.0704 | -34.1318 | -5.8738
3.0765 | -34.1322 | -5.885
3.0827 | -34.1325 | -5.8965
3.0888 | -34.1329 | -5.908
3.0949 | -34.1334 | -5.9196
3.101 | -34.1338 | -5.9313
3.1071 | -34.1341 | -5.9433
3.1132 | -34.1346 | -5.9554
3.1193 | -34.135 | -5.9671
3.1254 | -34.1354 | -5.9786
3.1315 | -34.1358 | -5.9908
3.1376 | -34.1361 | -6.0024
3.1437 | -34.1365 | -6.0143
3.1498 | -34.1369 | -6.0267
3.1559 | -34.1373 | -6.0384
3.162 | -34.1376 | -6.0507
3.1681 | -34.1379 | -6.0628
3.1742 | -34.1382 | -6.0747
3.1803 | -34.1386 | -6.0867
3.1864 | -34.139 | -6.0981
3.1925 | -34.1393 | -6.1099
3.1986 | -34.1396 | -6.1221
3.2047 | -34.1399 | -6.1343
3.2108 | -34.1403 | -6.1457
3.2169 | -34.1406 | -6.1576
3.223 | -34.1409 | -6.1694
3.2292 | -34.1412 | -6.1814
3.2353 | -34.1414 | -6.1924
3.2414 | -34.1417 | -6.2039
3.2475 | -34.142 | -6.2157
3.2536 | -34.1423 | -6.2272
3.2597 | -34.1426 | -6.2383
3.2658 | -34.1428 | -6.2495
3.2719 | -34.1431 | -6.2607
3.278 | -34.1434 | -6.2726
3.2841 | -34.1437 | -6.2838
3.2902 | -34.144 | -6.2954
3.2963 | -34.1443 | -6.3071
3.3024 | -34.1444 | -6.3183
3.3085 | -34.1447 | -6.33
3.3146 | -34.1449 | -6.3415
3.3207 | -34.1452 | -6.3531
3.3268 | -34.1455 | -6.3647
3.3329 | -34.1457 | -6.3757
3.339 | -34.1459 | -6.3866
3.3451 | -34.1461 | -6.398
3.3512 | -34.1463 | -6.4088
3.3573 | -34.1466 | -6.4202
3.3634 | -34.1468 | -6.4316
3.3696 | -34.147 | -6.4432
3.3757 | -34.1472 | -6.454
3.3818 | -34.1475 | -6.4651
3.3879 | -34.1477 | -6.4763
3.394 | -34.1479 | -6.4878
3.4001 | -34.1481 | -6.4988
3.4062 | -34.1482 | -6.5098
3.4123 | -34.1484 | -6.5205
3.4184 | -34.1487 | -6.532
3.4245 | -34.1489 | -6.5429
3.4306 | -34.149 | -6.5537
3.4367 | -34.1493 | -6.5643
3.4428 | -34.1495 | -6.5748
3.4489 | -34.1497 | -6.5859
3.455 | -34.1499 | -6.597
3.4611 | -34.1501 | -6.6084
3.4672 | -34.1503 | -6.6195
3.4733 | -34.1505 | -6.6307
3.4794 | -34.1507 | -6.6416
3.4855 | -34.1509 | -6.6521
3.4916 | -34.1511 | -6.6626
3.4977 | -34.1513 | -6.674
3.5038 | -34.1515 | -6.685
3.5099 | -34.1517 | -6.6954
3.5161 | -34.1519 | -6.706
3.5222 | -34.152 | -6.7164
3.5283 | -34.1522 | -6.7273
3.5344 | -34.1523 | -6.7383
3.5405 | -34.1526 | -6.7489
3.5466 | -34.1528 | -6.76
3.5527 | -34.1529 | -6.771
3.5588 | -34.153 | -6.7817
3.5649 | -34.1531 | -6.7924
3.571 | -34.1533 | -6.8032
3.5771 | -34.1535 | -6.8144
3.5832 | -34.1537 | -6.8242
3.5893 | -34.1539 | -6.835
3.5954 | -34.1541 | -6.8459
3.6015 | -34.1544 | -6.857
3.6076 | -34.1546 | -6.8676
3.6137 | -34.1548 | -6.8778
3.6198 | -34.1551 | -6.8881
3.6259 | -34.1553 | -6.8992
3.632 | -34.1555 | -6.9101
3.6381 | -34.1557 | -6.9209
3.6442 | -34.1559 | -6.9317
3.6503 | -34.1562 | -6.9419
3.6565 | -34.1564 | -6.9529
3.6626 | -34.1566 | -6.9636
3.6687 | -34.1569 | -6.9744
3.6748 | -34.1572 | -6.9849
3.6809 | -34.1575 | -6.9957
3.687 | -34.1577 | -7.0068
3.6931 | -34.158 | -7.0172
3.6992 | -34.1582 | -7.028
3.7053 | -34.1584 | -7.0387
3.7114 | -34.1586 | -7.0494
3.7175 | -34.1589 | -7.0603
3.7236 | -34.1592 | -7.0705
3.7297 | -34.1594 | -7.0815
3.7358 | -34.1597 | -7.0925
3.7419 | -34.1599 | -7.1035
3.748 | -34.1602 | -7.1141
3.7541 | -34.1604 | -7.1248
3.7602 | -34.1608 | -7.1356
3.7663 | -34.161 | -7.1461
3.7724 | -34.1613 | -7.157
3.7785 | -34.1616 | -7.1676
3.7846 | -34.1618 | -7.1782
3.7907 | -34.1622 | -7.1894
3.7969 | -34.1625 | -7.2003
3.803 | -34.1628 | -7.2107
3.8091 | -34.163 | -7.2214
3.8152 | -34.1634 | -7.2321
3.8213 | -34.1636 | -7.2432
3.8274 | -34.1639 | -7.2537
3.8335 | -34.1643 | -7.2647
3.8396 | -34.1646 | -7.2755
3.8457 | -34.165 | -7.2865
3.8518 | -34.1654 | -7.2971
3.8579 | -34.1657 | -7.3079
3.864 | -34.1661 | -7.3189
3.8701 | -34.1664 | -7.3301
3.8762 | -34.1667 | -7.3408
3.8823 | -34.167 | -7.3518
3.8884 | -34.1674 | -7.3627
3.8945 | -34.1678 | -7.3738
3.9006 | -34.1681 | -7.385
3.9067 | -34.1684 | -7.3963
3.9128 | -34.1687 | -7.4072
3.9189 | -34.1691 | -7.4179
3.925 | -34.1693 | -7.4289
3.9311 | -34.1697 | -7.4393
3.9372 | -34.17 | -7.4504
3.9434 | -34.1703 | -7.4613
3.9495 | -34.1707 | -7.4723
3.9556 | -34.1711 | -7.483
3.9617 | -34.1714 | -7.4935
3.9678 | -34.1718 | -7.5042
3.9739 | -34.1721 | -7.515
3.98 | -34.1724 | -7.5259
3.9861 | -34.1728 | -7.5368
3.9922 | -34.173 | -7.5476
3.9983 | -34.1734 | -7.5581
4.0044 | -34.1737 | -7.5685
4.0105 | -34.174 | -7.5795
4.0166 | -34.1744 | -7.5904
4.0227 | -34.1747 | -7.6012
4.0288 | -34.175 | -7.6125
4.0349 | -34.1753 | -7.6231
4.041 | -34.1756 | -7.634
4.0471 | -34.1758 | -7.6449
4.0532 | -34.1762 | -7.6554
4.0593 | -34.1766 | -7.6665
4.0654 | -34.1769 | -7.6771
4.0715 | -34.1772 | -7.6885
4.0776 | -34.1776 | -7.6993
4.0838 | -34.1779 | -7.7106
4.0899 | -34.1782 | -7.7212
4.096 | -34.1785 | -7.732
4.1021 | -34.1788 | -7.7425
4.1082 | -34.1791 | -7.7531
4.1143 | -34.1794 | -7.7636
4.1204 | -34.1798 | -7.7741
4.1265 | -34.1801 | -7.7846
4.1326 | -34.1804 | -7.7951
4.1387 | -34.1808 | -7.8057
4.1448 | -34.1811 | -7.8167
4.1509 | -34.1813 | -7.8276
4.157 | -34.1816 | -7.8387
4.1631 | -34.1819 | -7.8497
4.1692 | -34.1822 | -7.8604
4.1753 | -34.1826 | -7.8711
4.1814 | -34.1828 | -7.8815
4.1875 | -34.1831 | -7.8925
4.1936 | -34.1834 | -7.9033
4.1997 | -34.1837 | -7.9137
4.2058 | -34.184 | -7.9247
4.2119 | -34.1843 | -7.9353
4.218 | -34.1845 | -7.9457
4.2241 | -34.1848 | -7.9562
4.2303 | -34.1851 | -7.9666
4.2364 | -34.1854 | -7.9773
4.2425 | -34.1856 | -7.9881
4.2486 | -34.1859 | -7.9988
4.2547 | -34.1861 | -8.0092
4.2608 | -34.1863 | -8.0201
4.2669 | -34.1866 | -8.0304
4.273 | -34.1868 | -8.0412
4.2791 | -34.1871 | -8.0518
4.2852 | -34.1874 | -8.0628
4.2913 | -34.1876 | -8.0732
4.2974 | -34.1879 | -8.0841
4.3035 | -34.1882 | -8.0951
4.3096 | -34.1884 | -8.1059
4.3157 | -34.1886 | -8.1164
4.3218 | -34.1889 | -8.1268
4.3279 | -34.1892 | -8.1375
4.334 | -34.1894 | -8.148
4.3401 | -34.1896 | -8.1582
4.3462 | -34.1899 | -8.169
4.3523 | -34.1902 | -8.1792
4.3584 | -34.1904 | -8.1893
4.3645 | -34.1907 | -8.1998
4.3707 | -34.1909 | -8.2101
4.3768 | -34.1911 | -8.2206
4.3829 | -34.1914 | -8.2315
4.389 | -34.1917 | -8.242
4.3951 | -34.1919 | -8.2523
4.4012 | -34.1921 | -8.2632
4.4073 | -34.1923 | -8.2735
4.4134 | -34.1926 | -8.2839
4.4195 | -34.1929 | -8.2944
4.4256 | -34.1932 | -8.3047
4.4317 | -34.1934 | -8.3151
4.4378 | -34.1936 | -8.3252
4.4439 | -34.1939 | -8.3354
4.45 | -34.1941 | -8.3458
4.4561 | -34.1943 | -8.3563
4.4622 | -34.1945 | -8.3667
4.4683 | -34.1947 | -8.3771
4.4744 | -34.195 | -8.3879
4.4805 | -34.1952 | -8.3979
4.4866 | -34.1955 | -8.4083
4.4927 | -34.1957 | -8.419
4.4988 | -34.196 | -8.4294
4.5049 | -34.1962 | -8.4394
4.511 | -34.1965 | -8.4496
4.5172 | -34.1967 | -8.46
4.5233 | -34.197 | -8.4708
4.5294 | -34.1972 | -8.481
4.5355 | -34.1975 | -8.4913
4.5416 | -34.1977 | -8.5018
4.5477 | -34.198 | -8.5121
4.5538 | -34.1982 | -8.5221
4.5599 | -34.1985 | -8.5324
4.566 | -34.1987 | -8.5429
4.5721 | -34.199 | -8.5533
4.5782 | -34.1993 | -8.5638
4.5843 | -34.1995 | -8.5744
4.5904 | -34.1998 | -8.5849
4.5965 | -34.2001 | -8.5951
4.6026 | -34.2003 | -8.6059
4.6087 | -34.2005 | -8.6162
4.6148 | -34.2008 | -8.6269
4.6209 | -34.201 | -8.6373
4.627 | -34.2013 | -8.6479
4.6331 | -34.2015 | -8.6581
4.6392 | -34.2018 | -8.6685
4.6453 | -34.202 | -8.6789
4.6514 | -34.2023 | -8.6894
4.6576 | -34.2025 | -8.7002
4.6637 | -34.2027 | -8.7102
4.6698 | -34.2029 | -8.7204
4.6759 | -34.2032 | -8.7311
4.682 | -34.2034 | -8.7415
4.6881 | -34.2036 | -8.7519
4.6942 | -34.2039 | -8.7622
4.7003 | -34.2041 | -8.7727
4.7064 | -34.2043 | -8.7829
4.7125 | -34.2046 | -8.7936
4.7186 | -34.2048 | -8.8037
4.7247 | -34.2051 | -8.8142
4.7308 | -34.2054 | -8.8245
4.7369 | -34.2056 | -8.835
4.743 | -34.2059 | -8.8458
4.7491 | -34.2061 | -8.8566
4.7552 | -34.2064 | -8.8671
4.7613 | -34.2066 | -8.8776
4.7674 | -34.2068 | -8.8879
4.7735 | -34.207 | -8.8988
4.7796 | -34.2073 | -8.9095
4.7857 | -34.2076 | -8.92
4.7918 | -34.2079 | -8.9302
4.7979 | -34.2082 | -8.9405
4.8041 | -34.2084 | -8.951
4.8102 | -34.2087 | -8.9611
4.8163 | -34.2089 | -8.9715
4.8224 | -34.2092 | -8.982
4.8285 | -34.2095 | -8.9922
4.8346 | -34.2098 | -9.0024
4.8407 | -34.21 | -9.013
4.8468 | -34.2103 | -9.0237
4.8529 | -34.2105 | -9.0345
4.859 | -34.2108 | -9.0452
4.8651 | -34.211 | -9.0555
4.8712 | -34.2113 | -9.0658
4.8773 | -34.2115 | -9.0763
4.8834 | -34.2118 | -9.0868
4.8895 | -34.212 | -9.0971
4.8956 | -34.2122 | -9.1075
4.9017 | -34.2124 | -9.1178
4.9078 | -34.2127 | -9.1287
4.9139 | -34.213 | -9.1387
4.92 | -34.2133 | -9.1488
4.9261 | -34.2136 | -9.1588
4.9322 | -34.2139 | -9.1692
4.9383 | -34.2141 | -9.1797
4.9445 | -34.2144 | -9.1902
4.9506 | -34.2146 | -9.2005
4.9567 | -34.2149 | -9.2108
4.9628 | -34.2152 | -9.2212
4.9689 | -34.2155 | -9.2321
4.975 | -34.2157 | -9.2429
4.9811 | -34.216 | -9.253
4.9872 | -34.2163 | -9.2632
4.9933 | -34.2167 | -9.274
4.9994 | -34.217 | -9.2849
5.0055 | -34.2173 | -9.2954
5.0116 | -34.2176 | -9.3063
5.0177 | -34.2179 | -9.317
5.0238 | -34.2182 | -9.3279
5.0299 | -34.2186 | -9.3386
5.036 | -34.2189 | -9.3491
5.0421 | -34.2192 | -9.3595
5.0482 | -34.2195 | -9.37
5.0543 | -34.2198 | -9.3802
5.0604 | -34.2201 | -9.3908
5.0665 | -34.2204 | -9.401
5.0726 | -34.2207 | -9.4118
5.0787 | -34.221 | -9.4225
5.0848 | -34.2214 | -9.4332
5.091 | -34.2217 | -9.4434
5.0971 | -34.222 | -9.4536
5.1032 | -34.2223 | -9.4645
5.1093 | -34.2226 | -9.4751
5.1154 | -34.2229 | -9.4853
5.1215 | -34.2232 | -9.4956
5.1276 | -34.2235 | -9.5059
5.1337 | -34.2239 | -9.5162
5.1398 | -34.2241 | -9.5264
5.1459 | -34.2245 | -9.5372
5.152 | -34.2248 | -9.5479
5.1581 | -34.2251 | -9.5586
5.1642 | -34.2255 | -9.5691
5.1703 | -34.2258 | -9.5796
5.1764 | -34.2261 | -9.5899
5.1825 | -34.2264 | -9.6002
5.1886 | -34.2268 | -9.6103
5.1947 | -34.227 | -9.6208
5.2008 | -34.2274 | -9.6314
5.2069 | -34.2277 | -9.6421
5.213 | -34.2281 | -9.6524
5.2191 | -34.2284 | -9.663
5.2252 | -34.2288 | -9.6735
5.2314 | -34.2291 | -9.6843
5.2375 | -34.2294 | -9.6945
5.2436 | -34.2297 | -9.705
5.2497 | -34.23 | -9.7152
5.2558 | -34.2304 | -9.7261
5.2619 | -34.2307 | -9.7369
5.268 | -34.2311 | -9.7477
5.2741 | -34.2315 | -9.7583
5.2802 | -34.2318 | -9.7692
5.2863 | -34.2322 | -9.7797
5.2924 | -34.2326 | -9.7904
5.2985 | -34.2329 | -9.801
5.3046 | -34.2333 | -9.8116
5.3107 | -34.2336 | -9.8223
5.3168 | -34.234 | -9.8329
5.3229 | -34.2343 | -9.8434
5.329 | -34.2346 | -9.854
5.3351 | -34.235 | -9.8643
5.3412 | -34.2353 | -9.8748
5.3473 | -34.2357 | -9.8854
5.3534 | -34.2361 | -9.8961
5.3595 | -34.2364 | -9.9066
5.3656 | -34.2368 | -9.9171
5.3717 | -34.2371 | -9.928
5.3779 | -34.2375 | -9.9389
5.384 | -34.2378 | -9.9494
5.3901 | -34.2382 | -9.9603
5.3962 | -34.2386 | -9.9707
5.4023 | -34.2389 | -9.9814
5.4084 | -34.2393 | -9.9919
5.4145 | -34.2397 | -10.0022
5.4206 | -34.2401 | -10.0129
5.4267 | -34.2404 | -10.0238
5.4328 | -34.2408 | -10.0344
5.4389 | -34.2411 | -10.0447
5.445 | -34.2415 | -10.0556
5.4511 | -34.2418 | -10.0661
5.4572 | -34.2422 | -10.0767
5.4633 | -34.2426 | -10.0873
5.4694 | -34.2429 | -10.0979
5.4755 | -34.2432 | -10.1082
5.4816 | -34.2436 | -10.1189
5.4877 | -34.244 | -10.1294
5.4938 | -34.2443 | -10.1397
5.4999 | -34.2447 | -10.1501
5.506 | -34.2451 | -10.1606
5.5121 | -34.2454 | -10.1711
5.5183 | -34.2458 | -10.1814
5.5244 | -34.2461 | -10.1922
5.5305 | -34.2465 | -10.2026
5.5366 | -34.2469 | -10.213
5.5427 | -34.2472 | -10.2237
5.5488 | -34.2475 | -10.2344
5.5549 | -34.2479 | -10.2449
5.561 | -34.2482 | -10.2557
5.5671 | -34.2486 | -10.2661
5.5732 | -34.249 | -10.2769
5.5793 | -34.2493 | -10.2878
5.5854 | -34.2497 | -10.2982
5.5915 | -34.2501 | -10.3091
5.5976 | -34.2504 | -10.3198
5.6037 | -34.2508 | -10.3304
5.6098 | -34.2511 | -10.341
5.6159 | -34.2515 | -10.3513
5.622 | -34.2519 | -10.3619
5.6281 | -34.2522 | -10.3726
5.6342 | -34.2526 | -10.3834
5.6403 | -34.2529 | -10.3941
5.6464 | -34.2533 | -10.4048
5.6525 | -34.2536 | -10.4155
5.6586 | -34.254 | -10.4257
5.6648 | -34.2544 | -10.4364
5.6709 | -34.2548 | -10.4471
5.677 | -34.2551 | -10.4577
5.6831 | -34.2555 | -10.4684
5.6892 | -34.2558 | -10.4792
5.6953 | -34.2561 | -10.4897
5.7014 | -34.2565 | -10.5004
5.7075 | -34.2568 | -10.511
5.7136 | -34.2571 | -10.5215
5.7197 | -34.2575 | -10.532
5.7258 | -34.2578 | -10.5424
5.7319 | -34.2582 | -10.553
5.738 | -34.2585 | -10.5636
5.7441 | -34.2589 | -10.5741
5.7502 | -34.2593 | -10.5844
5.7563 | -34.2596 | -10.5949
5.7624 | -34.26 | -10.6058
5.7685 | -34.2603 | -10.6164
5.7746 | -34.2607 | -10.6272
5.7807 | -34.261 | -10.6378
5.7868 | -34.2613 | -10.6486
5.7929 | -34.2616 | -10.659
5.799 | -34.2619 | -10.6694
5.8052 | -34.2623 | -10.6798
5.8113 | -34.2626 | -10.6903
5.8174 | -34.2629 | -10.7007
5.8235 | -34.2633 | -10.7112
5.8296 | -34.2636 | -10.722
5.8357 | -34.2639 | -10.7324
5.8418 | -34.2643 | -10.7431
5.8479 | -34.2646 | -10.7537
5.854 | -34.2649 | -10.7644
5.8601 | -34.2653 | -10.7749
5.8662 | -34.2656 | -10.7856
5.8723 | -34.2659 | -10.7959
5.8784 | -34.2663 | -10.8063
5.8845 | -34.2666 | -10.8168
5.8906 | -34.2669 | -10.8275
5.8967 | -34.2672 | -10.8378
5.9028 | -34.2676 | -10.8481
5.9089 | -34.268 | -10.8585
5.915 | -34.2683 | -10.8688
5.9211 | -34.2686 | -10.8792
5.9272 | -34.269 | -10.8894
5.9333 | -34.2693 | -10.8999
5.9394 | -34.2697 | -10.9104
5.9455 | -34.27 | -10.9207
5.9517 | -34.2704 | -10.931
5.9578 | -34.2708 | -10.9416
5.9639 | -34.2711 | -10.9518
5.97 | -34.2715 | -10.9624
5.9761 | -34.2718 | -10.9727
5.9822 | -34.2722 | -10.9834
5.9883 | -34.2725 | -10.9939
5.9944 | -34.2728 | -11.0044
6.0005 | -34.2732 | -11.0151
6.0066 | -34.2735 | -11.0254
6.0127 | -34.2739 | -11.0355
6.0188 | -34.2742 | -11.0463
6.0249 | -34.2745 | -11.0568
6.031 | -34.2749 | -11.0671
6.0371 | -34.2752 | -11.0775
6.0432 | -34.2755 | -11.0879
6.0493 | -34.2759 | -11.0982
6.0554 | -34.2762 | -11.1085
6.0615 | -34.2766 | -11.119
6.0676 | -34.2769 | -11.1296
6.0737 | -34.2773 | -11.1402
6.0798 | -34.2776 | -11.1507
6.0859 | -34.278 | -11.1611
6.0921 | -34.2783 | -11.1712
6.0982 | -34.2786 | -11.1816
6.1043 | -34.279 | -11.1919
6.1104 | -34.2793 | -11.2025
6.1165 | -34.2797 | -11.2125
6.1226 | -34.28 | -11.2229
6.1287 | -34.2804 | -11.2332
6.1348 | -34.2807 | -11.2436
6.1409 | -34.2811 | -11.254
6.147 | -34.2814 | -11.2643
6.1531 | -34.2817 | -11.2747
6.1592 | -34.2821 | -11.2849
6.1653 | -34.2824 | -11.2952
6.1714 | -34.2827 | -11.3057
6.1775 | -34.2831 | -11.3162
6.1836 | -34.2834 | -11.3266
6.1897 | -34.2837 | -11.337
6.1958 | -34.2841 | -11.3477
6.2019 | -34.2844 | -11.3581
6.208 | -34.2848 | -11.3685
6.2141 | -34.2851 | -11.3789
6.2202 | -34.2855 | -11.3894
6.2263 | -34.2858 | -11.3998
6.2325 | -34.2862 | -11.4099
6.2386 | -34.2865 | -11.42
6.2447 | -34.2869 | -11.4306
6.2508 | -34.2872 | -11.441
6.2569 | -34.2876 | -11.4512
6.263 | -34.2879 | -11.4615
6.2691 | -34.2882 | -11.4718
6.2752 | -34.2885 | -11.4823
6.2813 | -34.2888 | -11.4926
6.2874 | -34.2892 | -11.5029
6.2935 | -34.2895 | -11.5132
6.2996 | -34.2898 | -11.5235
6.3057 | -34.2902 | -11.5339
6.3118 | -34.2905 | -11.5444
6.3179 | -34.2908 | -11.5547
6.324 | -34.2911 | -11.5651
6.3301 | -34.2915 | -11.5754
6.3362 | -34.2918 | -11.5857
6.3423 | -34.2922 | -11.5962
6.3484 | -34.2925 | -11.6066
6.3545 | -34.2928 | -11.6168
6.3606 | -34.2932 | -11.6268
6.3667 | -34.2934 | -11.6371
6.3728 | -34.2938 | -11.6474
6.379 | -34.2941 | -11.6578
6.3851 | -34.2945 | -11.6679
6.3912 | -34.2948 | -11.6781
6.3973 | -34.2951 | -11.6885
6.4034 | -34.2954 | -11.6987
6.4095 | -34.2957 | -11.709
6.4156 | -34.2961 | -11.7194
6.4217 | -34.2964 | -11.7296
6.4278 | -34.2967 | -11.74
6.4339 | -34.2971 | -11.7502
6.44 | -34.2975 | -11.7604
6.4461 | -34.2978 | -11.7709
6.4522 | -34.2981 | -11.781
6.4583 | -34.2984 | -11.7914
6.4644 | -34.2987 | -11.8018
6.4705 | -34.299 | -11.8118
6.4766 | -34.2994 | -11.8222
6.4827 | -34.2997 | -11.8325
6.4888 | -34.3 | -11.8426
6.4949 | -34.3003 | -11.8526
6.501 | -34.3007 | -11.863
6.5071 | -34.301 | -11.8731
6.5132 | -34.3014 | -11.8834
6.5194 | -34.3017 | -11.8938
6.5255 | -34.302 | -11.9038
6.5316 | -34.3023 | -11.9143
6.5377 | -34.3026 | -11.9246
6.5438 | -34.303 | -11.9347
6.5499 | -34.3033 | -11.9449
6.556 | -34.3036 | -11.9551
6.5621 | -34.3039 | -11.9656
6.5682 | -34.3043 | -11.976
6.5743 | -34.3047 | -11.9861
6.5804 | -34.305 | -11.9965
6.5865 | -34.3053 | -12.0068
6.5926 | -34.3056 | -12.017
6.5987 | -34.306 | -12.027
6.6048 | -34.3063 | -12.0373
6.6109 | -34.3066 | -12.0477
6.617 | -34.3069 | -12.0581
6.6231 | -34.3073 | -12.0683
6.6292 | -34.3076 | -12.0786
6.6353 | -34.3079 | -12.0892
6.6414 | -34.3083 | -12.0993
6.6475 | -34.3086 | -12.1096
6.6536 | -34.3089 | -12.1199
6.6597 | -34.3092 | -12.1299
6.6659 | -34.3096 | -12.1401
6.672 | -34.3099 | -12.1503
6.6781 | -34.3102 | -12.1606
6.6842 | -34.3106 | -12.1707
6.6903 | -34.311 | -12.1809
6.6964 | -34.3113 | -12.1911
6.7025 | -34.3116 | -12.2012
6.7086 | -34.3119 | -12.2115
6.7147 | -34.3123 | -12.2218
6.7208 | -34.3126 | -12.2321
6.7269 | -34.3129 | -12.2423
6.733 | -34.3132 | -12.2524
6.7391 | -34.3136 | -12.2626
6.7452 | -34.3139 | -12.2726
6.7513 | -34.3142 | -12.2826
6.7574 | -34.3146 | -12.2929
6.7635 | -34.3149 | -12.3029
6.7696 | -34.3153 | -12.3131
6.7757 | -34.3156 | -12.3233
6.7818 | -34.316 | -12.3335
6.7879 | -34.3163 | -12.3436
6.794 | -34.3167 | -12.3536
6.8001 | -34.3169 | -12.3639
6.8063 | -34.3173 | -12.374
6.8124 | -34.3176 | -12.3844
6.8185 | -34.3179 | -12.3948
6.8246 | -34.3182 | -12.4049
6.8307 | -34.3185 | -12.415
6.8368 | -34.3189 | -12.4255
6.8429 | -34.3192 | -12.4357
6.849 | -34.3196 | -12.446
6.8551 | -34.3199 | -12.4561
6.8612 | -34.3203 | -12.4662
6.8673 | -34.3206 | -12.4764
6.8734 | -34.3209 | -12.4867
6.8795 | -34.3213 | -12.4969
6.8856 | -34.3216 | -12.5071
6.8917 | -34.3219 | -12.5173
6.8978 | -34.3223 | -12.5273
6.9039 | -34.3226 | -12.5372
6.91 | -34.323 | -12.5474
6.9161 | -34.3233 | -12.5574
6.9222 | -34.3236 | -12.5676
6.9283 | -34.324 | -12.5775
6.9344 | -34.3244 | -12.5876
6.9405 | -34.3248 | -12.5977
6.9466 | -34.3251 | -12.608
6.9528 | -34.3255 | -12.6184
6.9589 | -34.3259 | -12.6287
6.965 | -34.3262 | -12.639
6.9711 | -34.3266 | -12.6493
6.9772 | -34.327 | -12.6594
6.9833 | -34.3273 | -12.6695
6.9894 | -34.3276 | -12.68
6.9955 | -34.328 | -12.6899
7.0016 | -34.3284 | -12.7001
7.0077 | -34.3287 | -12.7103
7.0138 | -34.3291 | -12.7205
7.0199 | -34.3294 | -12.7305
7.026 | -34.3297 | -12.7407
7.0321 | -34.3301 | -12.7506
7.0382 | -34.3305 | -12.7605
7.0443 | -34.3308 | -12.7707
7.0504 | -34.3312 | -12.7809
7.0565 | -34.3315 | -12.791
7.0626 | -34.3319 | -12.8009
7.0687 | -34.3323 | -12.8106
7.0748 | -34.3326 | -12.8205
7.0809 | -34.333 | -12.8306
7.087 | -34.3333 | -12.8409
7.0932 | -34.3337 | -12.8511
7.0993 | -34.334 | -12.8612
7.1054 | -34.3344 | -12.8712
7.1115 | -34.3348 | -12.881
7.1176 | -34.3351 | -12.8911
7.1237 | -34.3354 | -12.9012
7.1298 | -34.3358 | -12.9111
7.1359 | -34.3361 | -12.9212
7.142 | -34.3365 | -12.9311
7.1481 | -34.3368 | -12.941
7.1542 | -34.3372 | -12.9512
7.1603 | -34.3375 | -12.9613
7.1664 | -34.3379 | -12.9713
7.1725 | -34.3383 | -12.9814
7.1786 | -34.3387 | -12.9914
7.1847 | -34.339 | -13.0015
7.1908 | -34.3394 | -13.0118
7.1969 | -34.3397 | -13.022
7.203 | -34.3401 | -13.0318
7.2091 | -34.3405 | -13.0418
7.2152 | -34.3409 | -13.0518
7.2213 | -34.3412 | -13.0617
7.2274 | -34.3416 | -13.0718
7.2335 | -34.3419 | -13.0817
7.2397 | -34.3423 | -13.0916
7.2458 | -34.3427 | -13.1014
7.2519 | -34.3431 | -13.1115
7.258 | -34.3435 | -13.1217
7.2641 | -34.3438 | -13.1317
7.2702 | -34.3442 | -13.142
7.2763 | -34.3446 | -13.1521
7.2824 | -34.345 | -13.1622
7.2885 | -34.3453 | -13.1723
7.2946 | -34.3457 | -13.1823
7.3007 | -34.346 | -13.1924
7.3068 | -34.3463 | -13.2026
7.3129 | -34.3467 | -13.2127
7.319 | -34.3471 | -13.2228
7.3251 | -34.3475 | -13.2326
7.3312 | -34.3478 | -13.2427
7.3373 | -34.3482 | -13.2525
7.3434 | -34.3486 | -13.2624
7.3495 | -34.349 | -13.2722
7.3556 | -34.3494 | -13.2823
7.3617 | -34.3498 | -13.2925
7.3678 | -34.3501 | -13.3024
7.3739 | -34.3505 | -13.3126
7.3801 | -34.3509 | -13.3229
7.3862 | -34.3512 | -13.3329
7.3923 | -34.3516 | -13.343
7.3984 | -34.352 | -13.3528
7.4045 | -34.3524 | -13.3627
7.4106 | -34.3528 | -13.3727
7.4167 | -34.3531 | -13.3828
7.4228 | -34.3535 | -13.3929
7.4289 | -34.3539 | -13.4029
7.435 | -34.3543 | -13.4127
7.4411 | -34.3546 | -13.4227
7.4472 | -34.355 | -13.4327
7.4533 | -34.3554 | -13.4425
7.4594 | -34.3558 | -13.4527
7.4655 | -34.3562 | -13.4625
7.4716 | -34.3566 | -13.4728
7.4777 | -34.3569 | -13.4827
7.4838 | -34.3573 | -13.4925
7.4899 | -34.3577 | -13.5023
7.496 | -34.358 | -13.512
7.5021 | -34.3584 | -13.5219
7.5082 | -34.3588 | -13.532
7.5143 | -34.3592 | -13.542
7.5204 | -34.3595 | -13.5521
7.5266 | -34.3599 | -13.5619
7.5327 | -34.3603 | -13.5717
7.5388 | -34.3606 | -13.582
7.5449 | -34.361 | -13.5922
7.551 | -34.3614 | -13.602
7.5571 | -34.3618 | -13.6122
7.5632 | -34.3622 | -13.6222
7.5693 | -34.3626 | -13.6321
7.5754 | -34.363 | -13.6419
7.5815 | -34.3633 | -13.6519
7.5876 | -34.3637 | -13.6621
7.5937 | -34.3641 | -13.6722
7.5998 | -34.3645 | -13.6821
7.6059 | -34.3649 | -13.6923
7.612 | -34.3653 | -13.7023
7.6181 | -34.3657 | -13.7122
7.6242 | -34.3661 | -13.7224
7.6303 | -34.3665 | -13.7325
7.6364 | -34.3668 | -13.7427
7.6425 | -34.3672 | -13.7527
7.6486 | -34.3676 | -13.7627
7.6547 | -34.368 | -13.7726
7.6608 | -34.3683 | -13.7826
7.667 | -34.3687 | -13.7926
7.6731 | -34.3691 | -13.8027
7.6792 | -34.3695 | -13.8129
7.6853 | -34.3699 | -13.823
7.6914 | -34.3702 | -13.8331
7.6975 | -34.3706 | -13.8432
7.7036 | -34.371 | -13.8532
7.7097 | -34.3714 | -13.8634
7.7158 | -34.3718 | -13.8734
7.7219 | -34.3722 | -13.8832
7.728 | -34.3726 | -13.8933
7.7341 | -34.373 | -13.9033
7.7402 | -34.3734 | -13.9134
7.7463 | -34.3738 | -13.9235
7.7524 | -34.3741 | -13.9337
7.7585 | -34.3745 | -13.9437
7.7646 | -34.3749 | -13.9538
7.7707 | -34.3753 | -13.9639
7.7768 | -34.3756 | -13.974
7.7829 | -34.376 | -13.9843
7.789 | -34.3764 | -13.9946
7.7951 | -34.3768 | -14.0046
7.8012 | -34.3772 | -14.0146
7.8073 | -34.3776 | -14.0248
7.8135 | -34.378 | -14.0348
7.8196 | -34.3784 | -14.045
7.8257 | -34.3788 | -14.0552
7.8318 | -34.3791 | -14.0654
7.8379 | -34.3795 | -14.0753
7.844 | -34.3799 | -14.0853
7.8501 | -34.3803 | -14.0955
7.8562 | -34.3807 | -14.1058
7.8623 | -34.3811 | -14.1157
7.8684 | -34.3814 | -14.1259
7.8745 | -34.3818 | -14.136
7.8806 | -34.3822 | -14.1461
7.8867 | -34.3826 | -14.156
7.8928 | -34.383 | -14.1658
7.8989 | -34.3834 | -14.1759
7.905 | -34.3838 | -14.1861
7.9111 | -34.3841 | -14.1962
7.9172 | -34.3845 | -14.2062
7.9233 | -34.3849 | -14.2164
7.9294 | -34.3853 | -14.2265
7.9355 | -34.3856 | -14.2364
7.9416 | -34.386 | -14.2465
7.9477 | -34.3864 | -14.2564
7.9539 | -34.3867 | -14.2664
7.96 | -34.3871 | -14.2765
7.9661 | -34.3875 | -14.2865
7.9722 | -34.3879 | -14.2964
7.9783 | -34.3883 | -14.3067
7.9844 | -34.3887 | -14.3167
7.9905 | -34.3891 | -14.3269
7.9966 | -34.3894 | -14.3371
8.0027 | -34.3898 | -14.3472
8.0088 | -34.3902 | -14.3572
8.0149 | -34.3905 | -14.3673
8.021 | -34.3909 | -14.3774
8.0271 | -34.3913 | -14.3877
8.0332 | -34.3917 | -14.3976
8.0393 | -34.3921 | -14.4078
8.0454 | -34.3925 | -14.4182
8.0515 | -34.3929 | -14.4282
8.0576 | -34.3933 | -14.4385
8.0637 | -34.3937 | -14.4484
8.0698 | -34.394 | -14.4584
8.0759 | -34.3944 | -14.4684
8.082 | -34.3948 | -14.4784
8.0881 | -34.3952 | -14.4884
8.0942 | -34.3956 | -14.4986
8.1004 | -34.396 | -14.5087
8.1065 | -34.3964 | -14.5187
8.1126 | -34.3967 | -14.5288
8.1187 | -34.3971 | -14.5389
8.1248 | -34.3975 | -14.5488
8.1309 | -34.3979 | -14.5587
8.137 | -34.3983 | -14.5685
8.1431 | -34.3986 | -14.5789
8.1492 | -34.399 | -14.5889
8.1553 | -34.3994 | -14.599
8.1614 | -34.3998 | -14.6093
8.1675 | -34.4002 | -14.6193
8.1736 | -34.4005 | -14.6295
8.1797 | -34.4009 | -14.6397
8.1858 | -34.4013 | -14.6498
8.1919 | -34.4017 | -14.6599
8.198 | -34.402 | -14.6698
8.2041 | -34.4024 | -14.6799
8.2102 | -34.4028 | -14.6899
8.2163 | -34.4032 | -14.6998
8.2224 | -34.4036 | -14.7099
8.2285 | -34.4039 | -14.7197
8.2346 | -34.4043 | -14.7297
8.2408 | -34.4047 | -14.7396
8.2469 | -34.4051 | -14.7498
8.253 | -34.4055 | -14.7597
8.2591 | -34.4059 | -14.7698
8.2652 | -34.4063 | -14.7799
8.2713 | -34.4067 | -14.79
8.2774 | -34.4071 | -14.8002
8.2835 | -34.4075 | -14.8104
8.2896 | -34.4078 | -14.8203
8.2957 | -34.4082 | -14.8305
8.3018 | -34.4086 | -14.8406
8.3079 | -34.409 | -14.8505
8.314 | -34.4094 | -14.8606
8.3201 | -34.4098 | -14.8706
8.3262 | -34.4102 | -14.8807
8.3323 | -34.4106 | -14.8907
8.3384 | -34.411 | -14.9008
8.3445 | -34.4113 | -14.9108
8.3506 | -34.4117 | -14.9209
8.3567 | -34.4121 | -14.9311
8.3628 | -34.4125 | -14.9411
8.3689 | -34.4129 | -14.951
8.375 | -34.4132 | -14.9611
8.3812 | -34.4136 | -14.971
8.3873 | -34.414 | -14.981
8.3934 | -34.4144 | -14.9912
8.3995 | -34.4148 | -15.0011
8.4056 | -34.4152 | -15.0111
8.4117 | -34.4155 | -15.0212
8.4178 | -34.4159 | -15.0312
8.4239 | -34.4163 | -15.0411
8.43 | -34.4167 | -15.0512
8.4361 | -34.417 | -15.0613
8.4422 | -34.4174 | -15.0712
8.4483 | -34.4178 | -15.0811
8.4544 | -34.4182 | -15.0912
8.4605 | -34.4186 | -15.101
8.4666 | -34.419 | -15.1113
8.4727 | -34.4193 | -15.1213
8.4788 | -34.4197 | -15.1313
8.4849 | -34.4201 | -15.1413
8.491 | -34.4205 | -15.1513
8.4971 | -34.4209 | -15.1612
8.5032 | -34.4212 | -15.1711
8.5093 | -34.4217 | -15.1809
8.5154 | -34.4221 | -15.191
8.5215 | -34.4225 | -15.2009
8.5277 | -34.4228 | -15.2109
8.5338 | -34.4232 | -15.2208
8.5399 | -34.4236 | -15.2308
8.546 | -34.424 | -15.2407
8.5521 | -34.4244 | -15.2507
8.5582 | -34.4247 | -15.2607
8.5643 | -34.4251 | -15.2704
8.5704 | -34.4255 | -15.2802
8.5765 | -34.4259 | -15.2903
8.5826 | -34.4263 | -15.2998
8.5887 | -34.4267 | -15.3098
8.5948 | -34.4271 | -15.3198
8.6009 | -34.4275 | -15.3296
8.607 | -34.4279 | -15.3396
8.6131 | -34.4283 | -15.3496
8.6192 | -34.4287 | -15.3594
8.6253 | -34.4291 | -15.3693
8.6314 | -34.4295 | -15.3792
8.6375 | -34.4298 | -15.3892
8.6436 | -34.4302 | -15.3991
8.6497 | -34.4306 | -15.409
8.6558 | -34.431 | -15.4189
8.6619 | -34.4314 | -15.4289
8.6681 | -34.4318 | -15.439
8.6742 | -34.4322 | -15.4489
8.6803 | -34.4326 | -15.4588
8.6864 | -34.433 | -15.4684
8.6925 | -34.4334 | -15.4781
8.6986 | -34.4337 | -15.4881
8.7047 | -34.4341 | -15.4983
8.7108 | -34.4345 | -15.5082
8.7169 | -34.4349 | -15.5181
8.723 | -34.4353 | -15.5279
8.7291 | -34.4357 | -15.5376
8.7352 | -34.4361 | -15.5477
8.7413 | -34.4364 | -15.5577
8.7474 | -34.4368 | -15.5675
8.7535 | -34.4372 | -15.5774
8.7596 | -34.4376 | -15.5874
8.7657 | -34.438 | -15.5974
8.7718 | -34.4384 | -15.6074
8.7779 | -34.4388 | -15.6174
8.784 | -34.4392 | -15.6273
8.7901 | -34.4396 | -15.6373
8.7962 | -34.44 | -15.6469
8.8023 | -34.4404 | -15.6568
8.8084 | -34.4408 | -15.6666
8.8146 | -34.4412 | -15.6765
8.8207 | -34.4415 | -15.6863
8.8268 | -34.442 | -15.696
8.8329 | -34.4423 | -15.7059
8.839 | -34.4427 | -15.7158
8.8451 | -34.4431 | -15.7256
8.8512 | -34.4435 | -15.7353
8.8573 | -34.4439 | -15.7452
8.8634 | -34.4443 | -15.7549
8.8695 | -34.4447 | -15.765
8.8756 | -34.4451 | -15.775
8.8817 | -34.4455 | -15.7847
8.8878 | -34.4459 | -15.7948
8.8939 | -34.4463 | -15.8046
8.9 | -34.4467 | -15.8145
8.9061 | -34.4471 | -15.8243
8.9122 | -34.4475 | -15.8342
8.9183 | -34.4479 | -15.8441
8.9244 | -34.4483 | -15.8542
8.9305 | -34.4487 | -15.8638
8.9366 | -34.4491 | -15.8736
8.9427 | -34.4495 | -15.8836
8.9488 | -34.4499 | -15.8934
8.955 | -34.4503 | -15.9032
8.9611 | -34.4507 | -15.9132
8.9672 | -34.4512 | -15.9231
8.9733 | -34.4516 | -15.9328
8.9794 | -34.452 | -15.9425
8.9855 | -34.4524 | -15.9525
8.9916 | -34.4528 | -15.9625
8.9977 | -34.4532 | -15.9723
9.0038 | -34.4536 | -15.982
9.0099 | -34.454 | -15.9919
9.016 | -34.4544 | -16.0018
9.0221 | -34.4548 | -16.0114
9.0282 | -34.4552 | -16.021
9.0343 | -34.4556 | -16.0308
9.0404 | -34.4561 | -16.0407
9.0465 | -34.4565 | -16.0507
9.0526 | -34.4569 | -16.0604
9.0587 | -34.4573 | -16.0703
9.0648 | -34.4577 | -16.0799
9.0709 | -34.4581 | -16.0898
9.077 | -34.4586 | -16.0997
9.0831 | -34.4589 | -16.1096
9.0892 | -34.4594 | -16.1195
9.0953 | -34.4598 | -16.1294
9.1015 | -34.4602 | -16.1392
9.1076 | -34.4606 | -16.1489
9.1137 | -34.461 | -16.1587
9.1198 | -34.4615 | -16.1684
9.1259 | -34.4619 | -16.1782
9.132 | -34.4623 | -16.1883
9.1381 | -34.4627 | -16.1979
9.1442 | -34.4632 | -16.2075
9.1503 | -34.4636 | -16.2174
9.1564 | -34.464 | -16.2272
9.1625 | -34.4644 | -16.2371
9.1686 | -34.4648 | -16.2468
9.1747 | -34.4652 | -16.2566
9.1808 | -34.4657 | -16.2664
9.1869 | -34.4661 | -16.2762
9.193 | -34.4665 | -16.2861
9.1991 | -34.4669 | -16.2959
9.2052 | -34.4673 | -16.3057
9.2113 | -34.4677 | -16.3155
9.2174 | -34.4682 | -16.3254
9.2235 | -34.4686 | -16.3352
9.2296 | -34.4691 | -16.345
9.2357 | -34.4695 | -16.355
9.2419 | -34.4699 | -16.3647
9.248 | -34.4704 | -16.3745
9.2541 | -34.4708 | -16.3843
9.2602 | -34.4712 | -16.3943
9.2663 | -34.4717 | -16.4038
9.2724 | -34.4721 | -16.4136
9.2785 | -34.4725 | -16.4234
9.2846 | -34.4729 | -16.433
9.2907 | -34.4734 | -16.4428
9.2968 | -34.4738 | -16.4527
9.3029 | -34.4742 | -16.4627
9.309 | -34.4746 | -16.4724
9.3151 | -34.475 | -16.482
9.3212 | -34.4755 | -16.4917
9.3273 | -34.4759 | -16.5015
9.3334 | -34.4764 | -16.5111
9.3395 | -34.4768 | -16.5207
9.3456 | -34.4772 | -16.5304
9.3517 | -34.4776 | -16.5403
9.3578 | -34.478 | -16.55
9.3639 | -34.4785 | -16.5596
9.37 | -34.4789 | -16.5694
9.3761 | -34.4793 | -16.5794
9.3822 | -34.4798 | -16.5892
9.3884 | -34.4802 | -16.5988
9.3945 | -34.4806 | -16.6086
9.4006 | -34.4811 | -16.6184
9.4067 | -34.4815 | -16.6282
9.4128 | -34.482 | -16.638
9.4189 | -34.4824 | -16.648
9.425 | -34.4829 | -16.6577
9.4311 | -34.4833 | -16.6674
9.4372 | -34.4837 | -16.6771
9.4433 | -34.4842 | -16.6869
9.4494 | -34.4846 | -16.6965
9.4555 | -34.4851 | -16.7064
9.4616 | -34.4855 | -16.716
9.4677 | -34.4859 | -16.7258
9.4738 | -34.4864 | -16.7356
9.4799 | -34.4868 | -16.7454
9.486 | -34.4872 | -16.7551
9.4921 | -34.4876 | -16.7648
9.4982 | -34.4881 | -16.7744
9.5043 | -34.4885 | -16.7841
9.5104 | -34.489 | -16.7937
9.5165 | -34.4894 | -16.8035
9.5226 | -34.4898 | -16.8134
9.5288 | -34.4903 | -16.823
9.5349 | -34.4907 | -16.8328
9.541 | -34.4911 | -16.8423
9.5471 | -34.4915 | -16.852
9.5532 | -34.492 | -16.8619
9.5593 | -34.4924 | -16.8717
9.5654 | -34.4929 | -16.8814
9.5715 | -34.4933 | -16.8911
9.5776 | -34.4937 | -16.901
9.5837 | -34.4941 | -16.9106
9.5898 | -34.4946 | -16.9204
9.5959 | -34.495 | -16.9302
9.602 | -34.4954 | -16.9398
9.6081 | -34.4959 | -16.9495
9.6142 | -34.4963 | -16.9591
9.6203 | -34.4967 | -16.9689
9.6264 | -34.4971 | -16.9785
9.6325 | -34.4976 | -16.9883
9.6386 | -34.498 | -16.998
9.6447 | -34.4984 | -17.0077
9.6508 | -34.4989 | -17.0173
9.6569 | -34.4993 | -17.027
9.663 | -34.4998 | -17.0369
9.6691 | -34.5002 | -17.0466
9.6753 | -34.5007 | -17.0562
9.6814 | -34.5011 | -17.0658
9.6875 | -34.5015 | -17.0753
9.6936 | -34.5019 | -17.0849
9.6997 | -34.5024 | -17.0947
9.7058 | -34.5029 | -17.1044
9.7119 | -34.5033 | -17.1142
9.718 | -34.5038 | -17.124
9.7241 | -34.5042 | -17.1336
9.7302 | -34.5046 | -17.1433
9.7363 | -34.5051 | -17.1529
9.7424 | -34.5055 | -17.1625
9.7485 | -34.5059 | -17.1722
9.7546 | -34.5064 | -17.1817
9.7607 | -34.5068 | -17.1916
9.7668 | -34.5072 | -17.2014
9.7729 | -34.5077 | -17.211
9.779 | -34.5082 | -17.2205
9.7851 | -34.5086 | -17.2303
9.7912 | -34.5091 | -17.24
9.7973 | -34.5095 | -17.2498
9.8034 | -34.51 | -17.2592
9.8095 | -34.5104 | -17.2691
9.8157 | -34.5109 | -17.2787
9.8218 | -34.5113 | -17.2886
9.8279 | -34.5117 | -17.2981
9.834 | -34.5121 | -17.3079
9.8401 | -34.5126 | -17.3176
9.8462 | -34.5131 | -17.327
9.8523 | -34.5135 | -17.3366
9.8584 | -34.5139 | -17.3463
9.8645 | -34.5143 | -17.3561
9.8706 | -34.5148 | -17.3659
9.8767 | -34.5152 | -17.3755
9.8828 | -34.5156 | -17.385
9.8889 | -34.516 | -17.3944
9.895 | -34.5165 | -17.4041
9.9011 | -34.5169 | -17.4138
9.9072 | -34.5174 | -17.4234
9.9133 | -34.5178 | -17.4331
9.9194 | -34.5182 | -17.4426
9.9255 | -34.5186 | -17.4522
9.9316 | -34.5191 | -17.4619
9.9377 | -34.5195 | -17.4714
9.9438 | -34.5199 | -17.4811
9.9499 | -34.5204 | -17.491
9.956 | -34.5208 | -17.5006
9.9622 | -34.5212 | -17.5103
9.9683 | -34.5217 | -17.52
9.9744 | -34.5221 | -17.5298
9.9805 | -34.5226 | -17.5394
9.9866 | -34.523 | -17.549
9.9927 | -34.5234 | -17.5585
9.9988 | -34.5238 | -17.5683
10.0049 | -34.5243 | -17.5779
10.011 | -34.5247 | -17.5876
10.0171 | -34.5252 | -17.5972
10.0232 | -34.5256 | -17.6066
10.0293 | -34.5261 | -17.6164
10.0354 | -34.5266 | -17.6261
10.0415 | -34.527 | -17.6357
10.0476 | -34.5274 | -17.6455
10.0537 | -34.5279 | -17.6553
10.0598 | -34.5283 | -17.6648
10.0659 | -34.5287 | -17.6746
10.072 | -34.5292 | -17.684
10.0781 | -34.5296 | -17.6935
10.0842 | -34.5301 | -17.7032
10.0903 | -34.5305 | -17.7128
10.0964 | -34.5309 | -17.7221
10.1026 | -34.5314 | -17.7318
10.1087 | -34.5318 | -17.7414
10.1148 | -34.5322 | -17.751
10.1209 | -34.5326 | -17.7605
10.127 | -34.5331 | -17.7699
10.1331 | -34.5335 | -17.7794
10.1392 | -34.534 | -17.7889
10.1453 | -34.5344 | -17.7984
10.1514 | -34.5348 | -17.808
10.1575 | -34.5353 | -17.8177
10.1636 | -34.5357 | -17.8271
10.1697 | -34.5361 | -17.8368
10.1758 | -34.5366 | -17.8462
10.1819 | -34.537 | -17.8557
10.188 | -34.5374 | -17.8653
10.1941 | -34.5379 | -17.8749
10.2002 | -34.5383 | -17.8844
10.2063 | -34.5387 | -17.894
10.2124 | -34.5392 | -17.9037
10.2185 | -34.5396 | -17.9133
10.2246 | -34.5401 | -17.9228
10.2307 | -34.5405 | -17.9324
10.2368 | -34.5409 | -17.9417
10.2429 | -34.5413 | -17.9512
10.2491 | -34.5418 | -17.9609
10.2552 | -34.5422 | -17.9704
10.2613 | -34.5427 | -17.9801
10.2674 | -34.5431 | -17.9897
10.2735 | -34.5435 | -17.9995
10.2796 | -34.544 | -18.009
10.2857 | -34.5444 | -18.0188
10.2918 | -34.5449 | -18.0284
10.2979 | -34.5453 | -18.038
10.304 | -34.5458 | -18.0473
10.3101 | -34.5462 | -18.057
10.3162 | -34.5466 | -18.0667
10.3223 | -34.5471 | -18.0761
10.3284 | -34.5475 | -18.0856
10.3345 | -34.5479 | -18.0951
10.3406 | -34.5484 | -18.1045
10.3467 | -34.5488 | -18.114
10.3528 | -34.5492 | -18.1235
10.3589 | -34.5496 | -18.133
10.365 | -34.5501 | -18.1427
10.3711 | -34.5505 | -18.1523
10.3772 | -34.551 | -18.1619
10.3833 | -34.5514 | -18.1714
10.3895 | -34.5519 | -18.1809
10.3956 | -34.5523 | -18.1905
10.4017 | -34.5528 | -18.1999
10.4078 | -34.5532 | -18.2094
10.4139 | -34.5536 | -18.2189
10.42 | -34.5541 | -18.2286
10.4261 | -34.5545 | -18.2381
10.4322 | -34.555 | -18.2476
10.4383 | -34.5554 | -18.2572
10.4444 | -34.5558 | -18.2665
10.4505 | -34.5562 | -18.2761
10.4566 | -34.5567 | -18.2855
10.4627 | -34.5571 | -18.2949
10.4688 | -34.5575 | -18.3043
10.4749 | -34.558 | -18.3139
10.481 | -34.5584 | -18.3234
10.4871 | -34.5589 | -18.3328
10.4932 | -34.5593 | -18.3421
10.4993 | -34.5597 | -18.3517
10.5054 | -34.5602 | -18.3612
10.5115 | -34.5606 | -18.3708
10.5176 | -34.5611 | -18.3801
10.5237 | -34.5616 | -18.3897
10.5298 | -34.562 | -18.3991
10.536 | -34.5624 | -18.4087
10.5421 | -34.5628 | -18.418
10.5482 | -34.5633 | -18.4276
10.5543 | -34.5637 | -18.4369
10.5604 | -34.5642 | -18.4462
10.5665 | -34.5646 | -18.4556
10.5726 | -34.565 | -18.4651
10.5787 | -34.5655 | -18.4745
10.5848 | -34.5659 | -18.4837
10.5909 | -34.5664 | -18.4933
10.597 | -34.5668 | -18.5028
10.6031 | -34.5672 | -18.5124
10.6092 | -34.5677 | -18.5219
10.6153 | -34.5681 | -18.5315
10.6214 | -34.5686 | -18.5407
10.6275 | -34.569 | -18.5501
10.6336 | -34.5694 | -18.5596
10.6397 | -34.5699 | -18.5692
10.6458 | -34.5703 | -18.5786
10.6519 | -34.5707 | -18.5882
10.658 | -34.5712 | -18.5975
10.6641 | -34.5716 | -18.6071
10.6702 | -34.572 | -18.6166
10.6764 | -34.5725 | -18.6261
10.6825 | -34.5729 | -18.6354
10.6886 | -34.5734 | -18.6447
10.6947 | -34.5738 | -18.6541
10.7008 | -34.5743 | -18.6636
10.7069 | -34.5747 | -18.6731
10.713 | -34.5752 | -18.6826
10.7191 | -34.5756 | -18.692
10.7252 | -34.5761 | -18.7011
10.7313 | -34.5765 | -18.7105
10.7374 | -34.5769 | -18.72
10.7435 | -34.5774 | -18.7294
10.7496 | -34.5778 | -18.7386
10.7557 | -34.5782 | -18.7481
10.7618 | -34.5787 | -18.7572
10.7679 | -34.5791 | -18.7666
10.774 | -34.5796 | -18.7762
10.7801 | -34.58 | -18.7856
10.7862 | -34.5804 | -18.7948
10.7923 | -34.5809 | -18.8043
10.7984 | -34.5813 | -18.8136
10.8045 | -34.5818 | -18.8228
10.8106 | -34.5822 | -18.8323
10.8168 | -34.5826 | -18.8416
10.8229 | -34.5831 | -18.851
10.829 | -34.5835 | -18.8603
10.8351 | -34.584 | -18.8694
10.8412 | -34.5844 | -18.8787
10.8473 | -34.5849 | -18.8879
10.8534 | -34.5853 | -18.8973
10.8595 | -34.5858 | -18.9067
10.8656 | -34.5862 | -18.9158
10.8717 | -34.5867 | -18.9251
10.8778 | -34.5871 | -18.9346
10.8839 | -34.5875 | -18.9439
10.89 | -34.588 | -18.9533
10.8961 | -34.5885 | -18.9626
10.9022 | -34.5889 | -18.9719
10.9083 | -34.5893 | -18.9812
10.9144 | -34.5898 | -18.9904
10.9205 | -34.5902 | -18.9998
10.9266 | -34.5906 | -19.0091
10.9327 | -34.5911 | -19.0183
10.9388 | -34.5916 | -19.0275
10.9449 | -34.592 | -19.0368
10.951 | -34.5924 | -19.0461
10.9571 | -34.5929 | -19.0555
10.9633 | -34.5933 | -19.0649
10.9694 | -34.5938 | -19.0744
10.9755 | -34.5942 | -19.0837
10.9816 | -34.5947 | -19.0931
10.9877 | -34.5951 | -19.1024
10.9938 | -34.5956 | -19.1119
10.9999 | -34.596 | -19.1211
11.006 | -34.5965 | -19.1301
11.0121 | -34.5969 | -19.1395
11.0182 | -34.5973 | -19.1488
11.0243 | -34.5978 | -19.1581
11.0304 | -34.5982 | -19.1674
11.0365 | -34.5987 | -19.1767
11.0426 | -34.5992 | -19.1859
11.0487 | -34.5996 | -19.1953
11.0548 | -34.6001 | -19.2045
11.0609 | -34.6005 | -19.2138
11.067 | -34.601 | -19.2231
11.0731 | -34.6014 | -19.2323
11.0792 | -34.6019 | -19.2415
11.0853 | -34.6023 | -19.2508
11.0914 | -34.6028 | -19.2603
11.0975 | -34.6032 | -19.2695
11.1037 | -34.6036 | -19.2787
11.1098 | -34.6041 | -19.2882
11.1159 | -34.6046 | -19.2972
11.122 | -34.605 | -19.3064
11.1281 | -34.6054 | -19.3157
11.1342 | -34.6059 | -19.3251
11.1403 | -34.6063 | -19.3344
11.1464 | -34.6068 | -19.3437
11.1525 | -34.6072 | -19.3528
11.1586 | -34.6077 | -19.3622
11.1647 | -34.6081 | -19.3714
11.1708 | -34.6086 | -19.3806
11.1769 | -34.609 | -19.3898
11.183 | -34.6095 | -19.399
11.1891 | -34.6099 | -19.4082
11.1952 | -34.6104 | -19.4174
11.2013 | -34.6109 | -19.4268
11.2074 | -34.6113 | -19.436
11.2135 | -34.6117 | -19.4454
11.2196 | -34.6122 | -19.4545
11.2257 | -34.6127 | -19.4638
11.2318 | -34.6131 | -19.473
11.2379 | -34.6136 | -19.4823
11.244 | -34.6141 | -19.4917
11.2502 | -34.6146 | -19.5008
11.2563 | -34.615 | -19.51
11.2624 | -34.6155 | -19.5192
11.2685 | -34.6159 | -19.5284
11.2746 | -34.6164 | -19.5375
11.2807 | -34.6168 | -19.5469
11.2868 | -34.6173 | -19.5562
11.2929 | -34.6178 | -19.5652
11.299 | -34.6182 | -19.5745
11.3051 | -34.6187 | -19.5839
11.3112 | -34.6191 | -19.5933
11.3173 | -34.6196 | -19.6023
11.3234 | -34.6201 | -19.6113
11.3295 | -34.6205 | -19.6207
11.3356 | -34.621 | -19.63
11.3417 | -34.6214 | -19.6392
11.3478 | -34.6219 | -19.6485
11.3539 | -34.6223 | -19.6578
11.36 | -34.6227 | -19.6669
11.3661 | -34.6232 | -19.676
11.3722 | -34.6237 | -19.685
11.3783 | -34.6241 | -19.6944
11.3844 | -34.6246 | -19.7036
11.3906 | -34.625 | -19.7128
11.3967 | -34.6255 | -19.7219
11.4028 | -34.626 | -19.7312
11.4089 | -34.6264 | -19.7402
11.415 | -34.6268 | -19.7494
11.4211 | -34.6273 | -19.7587
11.4272 | -34.6277 | -19.7679
11.4333 | -34.6282 | -19.777
11.4394 | -34.6286 | -19.7861
11.4455 | -34.629 | -19.7952
11.4516 | -34.6295 | -19.8044
11.4577 | -34.6299 | -19.8137
11.4638 | -34.6304 | -19.8226
11.4699 | -34.6309 | -19.8319
11.476 | -34.6313 | -19.8413
11.4821 | -34.6318 | -19.8505
11.4882 | -34.6323 | -19.8598
11.4943 | -34.6327 | -19.8691
11.5004 | -34.6332 | -19.878
11.5065 | -34.6337 | -19.8872
11.5126 | -34.6341 | -19.8962
11.5187 | -34.6345 | -19.9055
11.5248 | -34.635 | -19.9148
11.5309 | -34.6355 | -19.9239
11.5371 | -34.6359 | -19.933
11.5432 | -34.6364 | -19.9422
11.5493 | -34.6368 | -19.9515
11.5554 | -34.6373 | -19.9607
11.5615 | -34.6377 | -19.97
11.5676 | -34.6382 | -19.9791
11.5737 | -34.6387 | -19.9881
11.5798 | -34.6391 | -19.9971
11.5859 | -34.6396 | -20.0063
11.592 | -34.6401 | -20.0154
11.5981 | -34.6405 | -20.0244
11.6042 | -34.641 | -20.0335
11.6103 | -34.6414 | -20.0425
11.6164 | -34.6419 | -20.0516
11.6225 | -34.6423 | -20.0607
11.6286 | -34.6428 | -20.0698
11.6347 | -34.6432 | -20.0791
11.6408 | -34.6437 | -20.0881
11.6469 | -34.6441 | -20.0973
11.653 | -34.6446 | -20.1063
11.6591 | -34.6451 | -20.1154
11.6652 | -34.6455 | -20.1244
11.6713 | -34.646 | -20.1335
11.6775 | -34.6465 | -20.1424
11.6836 | -34.6469 | -20.1515
11.6897 | -34.6474 | -20.1606
11.6958 | -34.6479 | -20.1698
11.7019 | -34.6483 | -20.1791
11.708 | -34.6488 | -20.1882
11.7141 | -34.6493 | -20.1972
11.7202 | -34.6497 | -20.2063
11.7263 | -34.6501 | -20.2155
11.7324 | -34.6506 | -20.2247
11.7385 | -34.6511 | -20.2339
11.7446 | -34.6515 | -20.2429
11.7507 | -34.652 | -20.2517
11.7568 | -34.6524 | -20.2609
11.7629 | -34.6529 | -20.2701
11.769 | -34.6533 | -20.2791
11.7751 | -34.6538 | -20.2882
11.7812 | -34.6543 | -20.2971
11.7873 | -34.6547 | -20.3062
11.7934 | -34.6552 | -20.3153
11.7995 | -34.6557 | -20.3242
11.8056 | -34.6561 | -20.3333
11.8117 | -34.6566 | -20.3423
11.8178 | -34.657 | -20.3513
11.824 | -34.6575 | -20.3603
11.8301 | -34.658 | -20.3693
11.8362 | -34.6584 | -20.3783
11.8423 | -34.6589 | -20.3874
11.8484 | -34.6593 | -20.3965
11.8545 | -34.6598 | -20.4054
11.8606 | -34.6603 | -20.4144
11.8667 | -34.6607 | -20.4236
11.8728 | -34.6612 | -20.4326
11.8789 | -34.6616 | -20.4416
11.885 | -34.6621 | -20.4507
11.8911 | -34.6626 | -20.4598
11.8972 | -34.6631 | -20.4688
11.9033 | -34.6635 | -20.4778
11.9094 | -34.6639 | -20.4868
11.9155 | -34.6644 | -20.4959
11.9216 | -34.6649 | -20.5049
11.9277 | -34.6654 | -20.5138
11.9338 | -34.6658 | -20.5229
11.9399 | -34.6663 | -20.5321
11.946 | -34.6668 | -20.5411
11.9521 | -34.6672 | -20.5501
11.9582 | -34.6677 | -20.5591
11.9644 | -34.6681 | -20.568
11.9705 | -34.6686 | -20.577
11.9766 | -34.6691 | -20.586
11.9827 | -34.6696 | -20.5949
11.9888 | -34.6701 | -20.6039
11.9949 | -34.6705 | -20.6126
12.001 | -34.671 | -20.6217
12.0071 | -34.6715 | -20.6306
12.0132 | -34.6719 | -20.6395
12.0193 | -34.6724 | -20.6487
12.0254 | -34.6728 | -20.6576
12.0315 | -34.6733 | -20.6665
12.0376 | -34.6737 | -20.6754
12.0437 | -34.6742 | -20.6844
12.0498 | -34.6747 | -20.6934
12.0559 | -34.6751 | -20.7023
12.062 | -34.6756 | -20.711
12.0681 | -34.6761 | -20.7201
12.0742 | -34.6766 | -20.7289
12.0803 | -34.677 | -20.738
12.0864 | -34.6775 | -20.7471
12.0925 | -34.6779 | -20.7561
12.0986 | -34.6784 | -20.7649
12.1047 | -34.6789 | -20.7739
12.1109 | -34.6794 | -20.7827
12.117 | -34.6798 | -20.7916
12.1231 | -34.6803 | -20.8005
12.1292 | -34.6808 | -20.8094
12.1353 | -34.6812 | -20.8183
12.1414 | -34.6817 | -20.8273
12.1475 | -34.6822 | -20.8363
12.1536 | -34.6826 | -20.8454
12.1597 | -34.6831 | -20.8544
12.1658 | -34.6836 | -20.8633
12.1719 | -34.6841 | -20.872
12.178 | -34.6845 | -20.8811
12.1841 | -34.685 | -20.8901
12.1902 | -34.6855 | -20.8991
12.1963 | -34.686 | -20.9082
12.2024 | -34.6865 | -20.917
12.2085 | -34.6869 | -20.9259
12.2146 | -34.6874 | -20.9348
12.2207 | -34.6878 | -20.9438
12.2268 | -34.6883 | -20.9528
12.2329 | -34.6888 | -20.9617
12.239 | -34.6893 | -20.9703
12.2451 | -34.6898 | -20.9793
12.2513 | -34.6902 | -20.9883
12.2574 | -34.6907 | -20.9972
12.2635 | -34.6912 | -21.0061
12.2696 | -34.6916 | -21.015
12.2757 | -34.6921 | -21.0236
12.2818 | -34.6926 | -21.0325
12.2879 | -34.6931 | -21.0415
12.294 | -34.6935 | -21.0503
12.3001 | -34.694 | -21.0592
12.3062 | -34.6945 | -21.0681
12.3123 | -34.695 | -21.0767
12.3184 | -34.6954 | -21.0855
12.3245 | -34.6959 | -21.0942
12.3306 | -34.6964 | -21.1031
12.3367 | -34.6968 | -21.112
12.3428 | -34.6973 | -21.1209
12.3489 | -34.6978 | -21.1295
12.355 | -34.6983 | -21.1386
12.3611 | -34.6987 | -21.1476
12.3672 | -34.6992 | -21.1564
12.3733 | -34.6997 | -21.1653
12.3794 | -34.7002 | -21.1739
12.3855 | -34.7006 | -21.1827
12.3916 | -34.7011 | -21.1916
12.3978 | -34.7016 | -21.2005
12.4039 | -34.7021 | -21.2095
12.41 | -34.7025 | -21.2184
12.4161 | -34.703 | -21.2269
12.4222 | -34.7035 | -21.2357
12.4283 | -34.704 | -21.2446
12.4344 | -34.7045 | -21.2535
12.4405 | -34.7049 | -21.2625
12.4466 | -34.7054 | -21.2713
12.4527 | -34.7059 | -21.28
12.4588 | -34.7063 | -21.2888
12.4649 | -34.7068 | -21.2977
12.471 | -34.7073 | -21.3067
12.4771 | -34.7078 | -21.3155
12.4832 | -34.7082 | -21.3243
12.4893 | -34.7087 | -21.333
12.4954 | -34.7092 | -21.3417
12.5015 | -34.7097 | -21.3503
12.5076 | -34.7102 | -21.3593
12.5137 | -34.7107 | -21.3681
12.5198 | -34.7111 | -21.3769
12.5259 | -34.7116 | -21.3858
12.532 | -34.7121 | -21.3946
12.5382 | -34.7126 | -21.4035
12.5443 | -34.713 | -21.4123
12.5504 | -34.7135 | -21.4212
12.5565 | -34.714 | -21.4298
12.5626 | -34.7145 | -21.4388
12.5687 | -34.715 | -21.4477
12.5748 | -34.7154 | -21.4566
12.5809 | -34.7159 | -21.4654
12.587 | -34.7164 | -21.4743
12.5931 | -34.7169 | -21.4829
12.5992 | -34.7174 | -21.4918
12.6053 | -34.7178 | -21.5005
12.6114 | -34.7183 | -21.5092
12.6175 | -34.7188 | -21.5181
12.6236 | -34.7193 | -21.5269
12.6297 | -34.7198 | -21.5357
12.6358 | -34.7202 | -21.5444
12.6419 | -34.7207 | -21.5532
12.648 | -34.7212 | -21.562
12.6541 | -34.7217 | -21.5709
12.6602 | -34.7222 | -21.5794
12.6663 | -34.7226 | -21.5882
12.6724 | -34.7231 | -21.5969
12.6785 | -34.7236 | -21.6057
12.6847 | -34.7241 | -21.6143
12.6908 | -34.7245 | -21.6231
12.6969 | -34.7251 | -21.6317
12.703 | -34.7255 | -21.6406
12.7091 | -34.726 | -21.6494
12.7152 | -34.7265 | -21.658
12.7213 | -34.727 | -21.6668
12.7274 | -34.7274 | -21.6756
12.7335 | -34.7279 | -21.6841
12.7396 | -34.7284 | -21.6927
12.7457 | -34.7289 | -21.7015
12.7518 | -34.7293 | -21.7102
12.7579 | -34.7298 | -21.7191
12.764 | -34.7303 | -21.7279
12.7701 | -34.7308 | -21.7367
12.7762 | -34.7313 | -21.7454
12.7823 | -34.7318 | -21.7542
12.7884 | -34.7322 | -21.763
12.7945 | -34.7327 | -21.772
12.8006 | -34.7332 | -21.7807
12.8067 | -34.7337 | -21.7895
12.8128 | -34.7342 | -21.7981
12.8189 | -34.7347 | -21.8069
12.8251 | -34.7352 | -21.8157
12.8312 | -34.7356 | -21.8246
12.8373 | -34.7361 | -21.8333
12.8434 | -34.7366 | -21.8421
12.8495 | -34.7371 | -21.8507
12.8556 | -34.7375 | -21.8594
12.8617 | -34.738 | -21.8681
12.8678 | -34.7385 | -21.8771
12.8739 | -34.739 | -21.8857
12.88 | -34.7394 | -21.8944
12.8861 | -34.7399 | -21.9031
12.8922 | -34.7404 | -21.9119
12.8983 | -34.7409 | -21.9207
12.9044 | -34.7414 | -21.9293
12.9105 | -34.7418 | -21.9379
12.9166 | -34.7423 | -21.9466
12.9227 | -34.7428 | -21.9554
12.9288 | -34.7432 | -21.9642
12.9349 | -34.7437 | -21.973
12.941 | -34.7442 | -21.9818
12.9471 | -34.7447 | -21.9906
12.9532 | -34.7452 | -21.9992
12.9593 | -34.7457 | -22.0081
12.9654 | -34.7462 | -22.017
12.9716 | -34.7467 | -22.0257
12.9777 | -34.7472 | -22.0345
12.9838 | -34.7477 | -22.043
12.9899 | -34.7481 | -22.0517
12.996 | -34.7486 | -22.0604
13.0021 | -34.7491 | -22.0692
13.0082 | -34.7496 | -22.0779
13.0143 | -34.7501 | -22.0863
13.0204 | -34.7506 | -22.095
13.0265 | -34.7511 | -22.1036
13.0326 | -34.7516 | -22.1124
13.0387 | -34.752 | -22.1212
13.0448 | -34.7525 | -22.13
13.0509 | -34.753 | -22.1388
13.057 | -34.7535 | -22.1474
13.0631 | -34.754 | -22.1559
13.0692 | -34.7545 | -22.1645
13.0753 | -34.755 | -22.1733
13.0814 | -34.7555 | -22.1821
13.0875 | -34.756 | -22.1908
13.0936 | -34.7565 | -22.1996
13.0997 | -34.757 | -22.2082
13.1058 | -34.7575 | -22.2168
13.112 | -34.758 | -22.2256
13.1181 | -34.7584 | -22.2343
13.1242 | -34.7589 | -22.2432
13.1303 | -34.7594 | -22.252
13.1364 | -34.7598 | -22.2609
13.1425 | -34.7603 | -22.2695
13.1486 | -34.7608 | -22.2784
13.1547 | -34.7613 | -22.287
13.1608 | -34.7617 | -22.2957
13.1669 | -34.7622 | -22.3042
13.173 | -34.7627 | -22.3131
13.1791 | -34.7632 | -22.3216
13.1852 | -34.7637 | -22.3302
13.1913 | -34.7642 | -22.3386
13.1974 | -34.7646 | -22.3474
13.2035 | -34.7651 | -22.356
13.2096 | -34.7656 | -22.3644
13.2157 | -34.7661 | -22.3731
13.2218 | -34.7666 | -22.3818
13.2279 | -34.7671 | -22.3906
13.234 | -34.7675 | -22.3992
13.2401 | -34.7681 | -22.4079
13.2462 | -34.7685 | -22.4166
13.2524 | -34.769 | -22.4255
13.2585 | -34.7695 | -22.4341
13.2646 | -34.77 | -22.4429
13.2707 | -34.7705 | -22.4517
13.2768 | -34.771 | -22.4606
13.2829 | -34.7715 | -22.4693
13.289 | -34.7719 | -22.4778
13.2951 | -34.7725 | -22.4862
13.3012 | -34.7729 | -22.495
13.3073 | -34.7735 | -22.5037
13.3134 | -34.7739 | -22.5124
13.3195 | -34.7744 | -22.521
13.3256 | -34.7749 | -22.5297
13.3317 | -34.7754 | -22.5383
13.3378 | -34.7759 | -22.547
13.3439 | -34.7764 | -22.5558
13.35 | -34.7769 | -22.5648
13.3561 | -34.7773 | -22.5735
13.3622 | -34.7778 | -22.5821
13.3683 | -34.7783 | -22.5905
13.3744 | -34.7788 | -22.5992
13.3805 | -34.7793 | -22.6078
13.3866 | -34.7798 | -22.6166
13.3927 | -34.7802 | -22.6252
13.3989 | -34.7807 | -22.6339
13.405 | -34.7812 | -22.6424
13.4111 | -34.7816 | -22.6511
13.4172 | -34.7821 | -22.6599
13.4233 | -34.7826 | -22.6686
13.4294 | -34.783 | -22.6775
13.4355 | -34.7835 | -22.6862
13.4416 | -34.784 | -22.6948
13.4477 | -34.7845 | -22.7034
13.4538 | -34.7851 | -22.7123
13.4599 | -34.7855 | -22.721
13.466 | -34.786 | -22.7298
13.4721 | -34.7865 | -22.7382
13.4782 | -34.787 | -22.7469
13.4843 | -34.7875 | -22.7557
13.4904 | -34.788 | -22.7644
13.4965 | -34.7884 | -22.7732
13.5026 | -34.7889 | -22.782
13.5087 | -34.7894 | -22.7906
13.5148 | -34.7899 | -22.7995
13.5209 | -34.7903 | -22.8083
13.527 | -34.7908 | -22.8169
13.5331 | -34.7913 | -22.8256
13.5393 | -34.7918 | -22.8343
13.5454 | -34.7923 | -22.843
13.5515 | -34.7928 | -22.8519
13.5576 | -34.7932 | -22.8605
13.5637 | -34.7937 | -22.8693
13.5698 | -34.7942 | -22.8781
13.5759 | -34.7947 | -22.8865
13.582 | -34.7952 | -22.8952
13.5881 | -34.7957 | -22.904
13.5942 | -34.7962 | -22.9128
13.6003 | -34.7967 | -22.9216
13.6064 | -34.7972 | -22.9306
13.6125 | -34.7976 | -22.9392
13.6186 | -34.7981 | -22.948
13.6247 | -34.7986 | -22.9568
13.6308 | -34.7991 | -22.9655
13.6369 | -34.7996 | -22.9743
13.643 | -34.8001 | -22.9832
13.6491 | -34.8005 | -22.9917
13.6552 | -34.801 | -23.0004
13.6613 | -34.8015 | -23.0091
13.6674 | -34.8019 | -23.0178
13.6735 | -34.8024 | -23.0265
13.6796 | -34.8029 | -23.0351
13.6858 | -34.8034 | -23.0437
13.6919 | -34.8039 | -23.0522
13.698 | -34.8044 | -23.0611
13.7041 | -34.8048 | -23.0697
13.7102 | -34.8054 | -23.0782
13.7163 | -34.8059 | -23.0868
13.7224 | -34.8063 | -23.0955
13.7285 | -34.8068 | -23.1042
13.7346 | -34.8073 | -23.1129
13.7407 | -34.8078 | -23.1216
13.7468 | -34.8083 | -23.1303
13.7529 | -34.8087 | -23.1386
13.759 | -34.8091 | -23.1473
13.7651 | -34.8096 | -23.156
13.7712 | -34.8101 | -23.1648
13.7773 | -34.8106 | -23.1733
13.7834 | -34.8111 | -23.1819
13.7895 | -34.8116 | -23.1904
13.7956 | -34.812 | -23.1991
13.8017 | -34.8124 | -23.2077
13.8078 | -34.8129 | -23.2162
13.8139 | -34.8134 | -23.2246
13.82 | -34.8138 | -23.2334
13.8262 | -34.8143 | -23.242
13.8323 | -34.8148 | -23.2508
13.8384 | -34.8153 | -23.2595
13.8445 | -34.8158 | -23.2681
13.8506 | -34.8163 | -23.2766
13.8567 | -34.8167 | -23.2854
13.8628 | -34.8172 | -23.294
13.8689 | -34.8177 | -23.3027
13.875 | -34.8182 | -23.3112
13.8811 | -34.8186 | -23.3199
13.8872 | -34.8191 | -23.3284
13.8933 | -34.8195 | -23.337
13.8994 | -34.82 | -23.3457
13.9055 | -34.8205 | -23.3544
13.9116 | -34.8211 | -23.363
13.9177 | -34.8215 | -23.3716
13.9238 | -34.822 | -23.3804
13.9299 | -34.8224 | -23.3891
13.936 | -34.8229 | -23.3979
13.9421 | -34.8234 | -23.4064
13.9482 | -34.8238 | -23.415
13.9543 | -34.8243 | -23.4238
13.9604 | -34.8248 | -23.4323
13.9665 | -34.8253 | -23.4408
13.9727 | -34.8258 | -23.4497
13.9788 | -34.8263 | -23.4584
13.9849 | -34.8268 | -23.4671
13.991 | -34.8273 | -23.4758
13.9971 | -34.8278 | -23.4842
14.0032 | -34.8282 | -23.4927
14.0093 | -34.8287 | -23.5013
14.0154 | -34.8292 | -23.5099
14.0215 | -34.8297 | -23.5182
14.0276 | -34.8302 | -23.5271
14.0337 | -34.8307 | -23.5353
14.0398 | -34.8312 | -23.5437
14.0459 | -34.8317 | -23.5525
14.052 | -34.8321 | -23.5613
14.0581 | -34.8326 | -23.5701
14.0642 | -34.833 | -23.5787
14.0703 | -34.8335 | -23.5873
14.0764 | -34.834 | -23.5957
14.0825 | -34.8345 | -23.6043
14.0886 | -34.835 | -23.6129
14.0947 | -34.8355 | -23.6215
14.1008 | -34.836 | -23.63
14.1069 | -34.8364 | -23.6386
14.1131 | -34.8369 | -23.6473
14.1192 | -34.8374 | -23.6558
14.1253 | -34.8379 | -23.6643
14.1314 | -34.8384 | -23.6729
14.1375 | -34.8389 | -23.6816
14.1436 | -34.8394 | -23.6901
14.1497 | -34.8399 | -23.6987
14.1558 | -34.8403 | -23.7074
14.1619 | -34.8408 | -23.7159
14.168 | -34.8413 | -23.7246
14.1741 | -34.8418 | -23.7331
14.1802 | -34.8423 | -23.7417
14.1863 | -34.8428 | -23.7505
14.1924 | -34.8433 | -23.7592
14.1985 | -34.8438 | -23.7677
14.2046 | -34.8442 | -23.7764
14.2107 | -34.8447 | -23.7847
14.2168 | -34.8452 | -23.7933
14.2229 | -34.8457 | -23.802
14.229 | -34.8462 | -23.8108
14.2351 | -34.8466 | -23.8194
14.2412 | -34.8471 | -23.8281
14.2473 | -34.8476 | -23.8366
14.2534 | -34.8481 | -23.845
14.2596 | -34.8485 | -23.8535
14.2657 | -34.849 | -23.8621
14.2718 | -34.8495 | -23.8707
14.2779 | -34.85 | -23.8793
14.284 | -34.8505 | -23.888
14.2901 | -34.851 | -23.8965
14.2962 | -34.8515 | -23.9052
14.3023 | -34.8519 | -23.9137
14.3084 | -34.8524 | -23.9222
14.3145 | -34.8529 | -23.9308
14.3206 | -34.8534 | -23.9395
14.3267 | -34.8539 | -23.9482
14.3328 | -34.8544 | -23.9566
14.3389 | -34.8549 | -23.9653
14.345 | -34.8553 | -23.974
14.3511 | -34.8558 | -23.9825
14.3572 | -34.8563 | -23.991
14.3633 | -34.8567 | -23.9998
14.3694 | -34.8572 | -24.0084
14.3755 | -34.8577 | -24.017
14.3816 | -34.8582 | -24.0257
14.3877 | -34.8587 | -24.034
14.3938 | -34.8591 | -24.0428
14.4 | -34.8596 | -24.0513
14.4061 | -34.8601 | -24.0598
14.4122 | -34.8606 | -24.0685
14.4183 | -34.861 | -24.0769
14.4244 | -34.8615 | -24.0856
14.4305 | -34.862 | -24.094
14.4366 | -34.8625 | -24.1027
14.4427 | -34.863 | -24.1114
14.4488 | -34.8634 | -24.1201
14.4549 | -34.8639 | -24.1283
14.461 | -34.8644 | -24.1367
14.4671 | -34.8649 | -24.1456
14.4732 | -34.8653 | -24.1542
14.4793 | -34.8658 | -24.1628
14.4854 | -34.8663 | -24.1714
14.4915 | -34.8668 | -24.1796
14.4976 | -34.8673 | -24.188
14.5037 | -34.8678 | -24.1966
14.5098 | -34.8683 | -24.2052
14.5159 | -34.8687 | -24.2138
14.522 | -34.8692 | -24.2222
14.5281 | -34.8697 | -24.2302
14.5342 | -34.8702 | -24.2388
14.5403 | -34.8707 | -24.2471
14.5465 | -34.8712 | -24.2557
14.5526 | -34.8717 | -24.2643
14.5587 | -34.8721 | -24.2729
14.5648 | -34.8726 | -24.2815
14.5709 | -34.8731 | -24.2901
14.577 | -34.8736 | -24.2986
14.5831 | -34.8741 | -24.3072
14.5892 | -34.8746 | -24.3157
14.5953 | -34.875 | -24.3239
14.6014 | -34.8755 | -24.3323
14.6075 | -34.8759 | -24.3407
14.6136 | -34.8764 | -24.3492
14.6197 | -34.8769 | -24.3577
14.6258 | -34.8773 | -24.3664
14.6319 | -34.8778 | -24.3749
14.638 | -34.8782 | -24.3832
14.6441 | -34.8787 | -24.3914
14.6502 | -34.8792 | -24.4001
14.6563 | -34.8798 | -24.4084
14.6624 | -34.8802 | -24.4169
14.6685 | -34.8807 | -24.4251
14.6746 | -34.8811 | -24.4338
14.6807 | -34.8816 | -24.4421
14.6869 | -34.8821 | -24.4507
14.693 | -34.8826 | -24.4592
14.6991 | -34.8831 | -24.4675
14.7052 | -34.8835 | -24.4758
14.7113 | -34.884 | -24.4843
14.7174 | -34.8845 | -24.4929
14.7235 | -34.885 | -24.5011
14.7296 | -34.8855 | -24.5098
14.7357 | -34.886 | -24.5183
14.7418 | -34.8865 | -24.5266
14.7479 | -34.887 | -24.535
14.754 | -34.8874 | -24.5433
14.7601 | -34.8879 | -24.5515
14.7662 | -34.8884 | -24.5598
14.7723 | -34.8888 | -24.5682
14.7784 | -34.8893 | -24.5765
14.7845 | -34.8898 | -24.5848
14.7906 | -34.8902 | -24.5933
14.7967 | -34.8907 | -24.6016
14.8028 | -34.8912 | -24.6099
14.8089 | -34.8916 | -24.6182
14.815 | -34.8921 | -24.6269
14.8211 | -34.8926 | -24.635
14.8272 | -34.893 | -24.6435
14.8334 | -34.8935 | -24.652
14.8395 | -34.894 | -24.6606
14.8456 | -34.8945 | -24.6692
14.8517 | -34.8949 | -24.6773
14.8578 | -34.8954 | -24.6856
14.8639 | -34.8959 | -24.694
14.87 | -34.8964 | -24.7026
14.8761 | -34.8969 | -24.7107
14.8822 | -34.8974 | -24.7193
14.8883 | -34.8979 | -24.7277
14.8944 | -34.8984 | -24.7358
14.9005 | -34.8989 | -24.7441
14.9066 | -34.8994 | -24.7524
14.9127 | -34.8998 | -24.7607
14.9188 | -34.9003 | -24.7692
14.9249 | -34.9008 | -24.7776
14.931 | -34.9013 | -24.7857
14.9371 | -34.9017 | -24.7941
14.9432 | -34.9022 | -24.8026
14.9493 | -34.9027 | -24.8105
14.9554 | -34.9031 | -24.8188
14.9615 | -34.9036 | -24.8275
14.9676 | -34.904 | -24.836
14.9738 | -34.9045 | -24.8444
14.9799 | -34.905 | -24.853
14.986 | -34.9056 | -24.861
14.9921 | -34.906 | -24.8694
14.9982 | -34.9065 | -24.8777
15.0043 | -34.907 | -24.8857
15.0104 | -34.9074 | -24.8943
15.0165 | -34.9079 | -24.9027
15.0226 | -34.9084 | -24.9103
15.0287 | -34.9088 | -24.919
15.0348 | -34.9092 | -24.9273
15.0409 | -34.9097 | -24.9355
15.047 | -34.9102 | -24.9439
15.0531 | -34.9108 | -24.9518
15.0592 | -34.9112 | -24.9604
15.0653 | -34.9117 | -24.9686
15.0714 | -34.9122 | -24.9772
15.0775 | -34.9127 | -24.9855
15.0836 | -34.9132 | -24.9934
15.0897 | -34.9138 | -25.0014
15.0958 | -34.9143 | -25.0099
15.1019 | -34.9148 | -25.0179
15.108 | -34.9152 | -25.0262
15.1141 | -34.9157 | -25.0344
15.1203 | -34.9162 | -25.043
15.1264 | -34.9166 | -25.0511
15.1325 | -34.9171 | -25.0594
15.1386 | -34.9176 | -25.0677
15.1447 | -34.918 | -25.0757
15.1508 | -34.9186 | -25.0844
15.1569 | -34.9191 | -25.0926
15.163 | -34.9196 | -25.1004
15.1691 | -34.92 | -25.1084
15.1752 | -34.9205 | -25.1164
15.1813 | -34.921 | -25.1247
15.1874 | -34.9214 | -25.1327
15.1935 | -34.922 | -25.1404
15.1996 | -34.9225 | -25.1482
15.2057 | -34.9229 | -25.1562
15.2118 | -34.9235 | -25.1643
15.2179 | -34.9239 | -25.1727
15.224 | -34.9244 | -25.181
15.2301 | -34.9249 | -25.1889
15.2362 | -34.9254 | -25.1975
15.2423 | -34.9259 | -25.2056
15.2484 | -34.9264 | -25.2137
15.2545 | -34.9268 | -25.222
15.2607 | -34.9273 | -25.2301
15.2668 | -34.9278 | -25.2377
15.2729 | -34.9283 | -25.2458
15.279 | -34.9288 | -25.254
15.2851 | -34.9293 | -25.2623
15.2912 | -34.9298 | -25.2707
15.2973 | -34.9303 | -25.279
15.3034 | -34.9308 | -25.287
15.3095 | -34.9313 | -25.2953
15.3156 | -34.9317 | -25.3036
15.3217 | -34.9322 | -25.3118
15.3278 | -34.9327 | -25.32
15.3339 | -34.9333 | -25.3276
15.34 | -34.9338 | -25.3358
15.3461 | -34.9343 | -25.3439
15.3522 | -34.9348 | -25.3521
15.3583 | -34.9353 | -25.3604
15.3644 | -34.9358 | -25.3686
15.3705 | -34.9363 | -25.3762
15.3766 | -34.9367 | -25.3843
15.3827 | -34.9373 | -25.3922
15.3888 | -34.9377 | -25.4007
15.3949 | -34.9382 | -25.4088
15.401 | -34.9387 | -25.4169
15.4072 | -34.9393 | -25.4249
15.4133 | -34.9397 | -25.4329
15.4194 | -34.9403 | -25.441
15.4255 | -34.9407 | -25.449
15.4316 | -34.9412 | -25.4573
15.4377 | -34.9417 | -25.4653
15.4438 | -34.9422 | -25.4734
15.4499 | -34.9427 | -25.4815
15.456 | -34.9432 | -25.4892
15.4621 | -34.9437 | -25.4972
15.4682 | -34.9442 | -25.5055
15.4743 | -34.9447 | -25.5136
15.4804 | -34.9452 | -25.522
15.4865 | -34.9457 | -25.5298
15.4926 | -34.9462 | -25.5381
15.4987 | -34.9467 | -25.5465
15.5048 | -34.9472 | -25.5545
15.5109 | -34.9477 | -25.5623
15.517 | -34.9482 | -25.5704
15.5231 | -34.9487 | -25.5783
15.5292 | -34.9492 | -25.5862
15.5353 | -34.9496 | -25.5943
15.5414 | -34.9501 | -25.6022
15.5476 | -34.9506 | -25.61
15.5537 | -34.9512 | -25.6179
15.5598 | -34.9517 | -25.6262
15.5659 | -34.9521 | -25.6342
15.572 | -34.9527 | -25.6421
15.5781 | -34.9532 | -25.6502
15.5842 | -34.9536 | -25.6582
15.5903 | -34.9541 | -25.6662
15.5964 | -34.9546 | -25.674
15.6025 | -34.9551 | -25.6822
15.6086 | -34.9556 | -25.6903
15.6147 | -34.9561 | -25.6983
15.6208 | -34.9566 | -25.7061
15.6269 | -34.9571 | -25.714
15.633 | -34.9576 | -25.7222
15.6391 | -34.9581 | -25.7303
15.6452 | -34.9586 | -25.738
15.6513 | -34.9591 | -25.7454
15.6574 | -34.9596 | -25.7537
15.6635 | -34.9601 | -25.7616
15.6696 | -34.9606 | -25.7698
15.6757 | -34.9611 | -25.7779
15.6818 | -34.9616 | -25.7858
15.688 | -34.9621 | -25.7936
15.6941 | -34.9626 | -25.8014
15.7002 | -34.9631 | -25.8095
15.7063 | -34.9635 | -25.8177
15.7124 | -34.964 | -25.8255
15.7185 | -34.9645 | -25.8336
15.7246 | -34.9651 | -25.8415
15.7307 | -34.9655 | -25.8493
15.7368 | -34.966 | -25.8571
15.7429 | -34.9665 | -25.8649
15.749 | -34.967 | -25.8731
15.7551 | -34.9675 | -25.8811
15.7612 | -34.9681 | -25.889
15.7673 | -34.9686 | -25.897
15.7734 | -34.9691 | -25.9048
15.7795 | -34.9696 | -25.9127
15.7856 | -34.97 | -25.9203
15.7917 | -34.9706 | -25.9275
15.7978 | -34.9711 | -25.9354
15.8039 | -34.9716 | -25.9433
15.81 | -34.9721 | -25.9512
15.8161 | -34.9726 | -25.959
15.8222 | -34.9731 | -25.9669
15.8283 | -34.9736 | -25.9747
15.8345 | -34.9741 | -25.9826
15.8406 | -34.9746 | -25.9904
15.8467 | -34.9751 | -25.9983
15.8528 | -34.9756 | -26.0061
15.8589 | -34.9761 | -26.0142
15.865 | -34.9767 | -26.0218
15.8711 | -34.9772 | -26.0296
15.8772 | -34.9777 | -26.0374
15.8833 | -34.9782 | -26.0454
15.8894 | -34.9787 | -26.053
15.8955 | -34.9792 | -26.0607
15.9016 | -34.9797 | -26.0685
15.9077 | -34.9802 | -26.0766
15.9138 | -34.9807 | -26.0843
15.9199 | -34.9812 | -26.092
15.926 | -34.9817 | -26.0999
15.9321 | -34.9822 | -26.1073
15.9382 | -34.9827 | -26.1149
15.9443 | -34.9832 | -26.1227
15.9504 | -34.9837 | -26.1305
15.9565 | -34.9842 | -26.1383
15.9626 | -34.9847 | -26.1464
15.9687 | -34.9852 | -26.1541
15.9749 | -34.9857 | -26.1618
15.981 | -34.9862 | -26.1696
15.9871 | -34.9867 | -26.1775
15.9932 | -34.9872 | -26.1852
15.9993 | -34.9877 | -26.1932
16.0054 | -34.9882 | -26.2007
16.0115 | -34.9887 | -26.2084
16.0176 | -34.9892 | -26.2163
16.0237 | -34.9897 | -26.2245
16.0298 | -34.9902 | -26.2323
16.0359 | -34.9907 | -26.2399
16.042 | -34.9912 | -26.2476
16.0481 | -34.9917 | -26.2553
16.0542 | -34.9922 | -26.2634
16.0603 | -34.9927 | -26.2711
16.0664 | -34.9932 | -26.2787
16.0725 | -34.9938 | -26.2862
16.0786 | -34.9942 | -26.2938
16.0847 | -34.9947 | -26.3016
16.0908 | -34.9952 | -26.3095
16.0969 | -34.9957 | -26.3172
16.103 | -34.9962 | -26.3249
16.1091 | -34.9968 | -26.3327
16.1152 | -34.9973 | -26.3405
16.1214 | -34.9978 | -26.3482
16.1275 | -34.9983 | -26.3559
16.1336 | -34.9988 | -26.3636
16.1397 | -34.9994 | -26.3712
16.1458 | -34.9999 | -26.3787
16.1519 | -35.0004 | -26.3866
16.158 | -35.0009 | -26.3943
16.1641 | -35.0014 | -26.4023
16.1702 | -35.0019 | -26.4101
16.1763 | -35.0024 | -26.4178
16.1824 | -35.0029 | -26.4253
16.1885 | -35.0034 | -26.4329
16.1946 | -35.0039 | -26.4405
16.2007 | -35.0044 | -26.4483
16.2068 | -35.0049 | -26.456
16.2129 | -35.0054 | -26.4638
16.219 | -35.006 | -26.471
16.2251 | -35.0065 | -26.4785
16.2312 | -35.007 | -26.4863
16.2373 | -35.0075 | -26.4939
16.2434 | -35.008 | -26.5018
16.2495 | -35.0086 | -26.5092
16.2556 | -35.0091 | -26.517
16.2618 | -35.0096 | -26.5247
16.2679 | -35.0102 | -26.5324
16.274 | -35.0107 | -26.5399
16.2801 | -35.0112 | -26.5475
16.2862 | -35.0117 | -26.5551
16.2923 | -35.0122 | -26.563
16.2984 | -35.0127 | -26.5707
16.3045 | -35.0132 | -26.5784
16.3106 | -35.0137 | -26.5861
16.3167 | -35.0143 | -26.5937
16.3228 | -35.0148 | -26.6011
16.3289 | -35.0153 | -26.6087
16.335 | -35.0158 | -26.6162
16.3411 | -35.0163 | -26.6238
16.3472 | -35.0169 | -26.6316
16.3533 | -35.0174 | -26.6394
16.3594 | -35.0179 | -26.6468
16.3655 | -35.0184 | -26.6545
16.3716 | -35.0189 | -26.6619
16.3777 | -35.0194 | -26.6697
16.3838 | -35.0199 | -26.6773
16.3899 | -35.0204 | -26.6849
16.396 | -35.0209 | -26.6926
16.4021 | -35.0215 | -26.7002
16.4083 | -35.022 | -26.7078
16.4144 | -35.0225 | -26.7155
16.4205 | -35.023 | -26.7231
16.4266 | -35.0235 | -26.7307
16.4327 | -35.024 | -26.7383
16.4388 | -35.0245 | -26.7459
16.4449 | -35.025 | -26.7535
16.451 | -35.0255 | -26.761
16.4571 | -35.026 | -26.7689
16.4632 | -35.0265 | -26.7763
16.4693 | -35.027 | -26.7837
16.4754 | -35.0275 | -26.7914
16.4815 | -35.028 | -26.7991
16.4876 | -35.0286 | -26.8067
16.4937 | -35.0291 | -26.8143
16.4998 | -35.0296 | -26.8219
16.5059 | -35.0301 | -26.8295
16.512 | -35.0306 | -26.8373
16.5181 | -35.0312 | -26.8448
16.5242 | -35.0317 | -26.8525
16.5303 | -35.0322 | -26.8599
16.5364 | -35.0327 | -26.8673
16.5425 | -35.0332 | -26.8749
16.5487 | -35.0337 | -26.8825
16.5548 | -35.0342 | -26.8902
16.5609 | -35.0347 | -26.8979
16.567 | -35.0353 | -26.9054
16.5731 | -35.0358 | -26.9131
16.5792 | -35.0363 | -26.9207
16.5853 | -35.0368 | -26.9283
16.5914 | -35.0373 | -26.9359
16.5975 | -35.0378 | -26.9436
16.6036 | -35.0384 | -26.9509
16.6097 | -35.0389 | -26.9587
16.6158 | -35.0394 | -26.9661
16.6219 | -35.0399 | -26.9736
16.628 | -35.0404 | -26.9813
16.6341 | -35.0409 | -26.989
16.6402 | -35.0415 | -26.9963
16.6463 | -35.0419 | -27.0038
16.6524 | -35.0424 | -27.0113
16.6585 | -35.043 | -27.0189
16.6646 | -35.0435 | -27.0264
16.6707 | -35.044 | -27.0337
16.6768 | -35.0445 | -27.0412
16.6829 | -35.045 | -27.0487
16.689 | -35.0456 | -27.0563
16.6952 | -35.046 | -27.0639
16.7013 | -35.0465 | -27.0714
16.7074 | -35.047 | -27.0788
16.7135 | -35.0475 | -27.0863
16.7196 | -35.048 | -27.094
16.7257 | -35.0486 | -27.1015
16.7318 | -35.0491 | -27.109
16.7379 | -35.0496 | -27.1166
16.744 | -35.0501 | -27.124
16.7501 | -35.0506 | -27.1316
16.7562 | -35.0511 | -27.1392
16.7623 | -35.0517 | -27.1467
16.7684 | -35.0521 | -27.1544
16.7745 | -35.0526 | -27.1619
16.7806 | -35.0532 | -27.1691
16.7867 | -35.0537 | -27.1767
16.7928 | -35.0542 | -27.1842
16.7989 | -35.0547 | -27.1916
16.805 | -35.0553 | -27.1992
16.8111 | -35.0558 | -27.2066
16.8172 | -35.0563 | -27.2142
16.8233 | -35.0568 | -27.2219
16.8294 | -35.0573 | -27.2294
16.8356 | -35.0578 | -27.237
16.8417 | -35.0583 | -27.2445
16.8478 | -35.0589 | -27.2518
16.8539 | -35.0594 | -27.2593
16.86 | -35.0599 | -27.2669
16.8661 | -35.0604 | -27.2744
16.8722 | -35.0609 | -27.282
16.8783 | -35.0614 | -27.2896
16.8844 | -35.062 | -27.297
16.8905 | -35.0625 | -27.3045
16.8966 | -35.063 | -27.3119
16.9027 | -35.0635 | -27.3194
16.9088 | -35.064 | -27.3271
16.9149 | -35.0645 | -27.3345
16.921 | -35.065 | -27.3417
16.9271 | -35.0655 | -27.3494
16.9332 | -35.0661 | -27.3569
16.9393 | -35.0666 | -27.3643
16.9454 | -35.0671 | -27.3719
16.9515 | -35.0676 | -27.3797
16.9576 | -35.0681 | -27.3869
16.9637 | -35.0686 | -27.3942
16.9698 | -35.0692 | -27.4017
16.9759 | -35.0696 | -27.4093
16.9821 | -35.0702 | -27.417
16.9882 | -35.0708 | -27.4245
16.9943 | -35.0713 | -27.4319
17.0004 | -35.0718 | -27.4393
17.0065 | -35.0723 | -27.4468
17.0126 | -35.0728 | -27.4544
17.0187 | -35.0734 | -27.462
17.0248 | -35.0739 | -27.4694
17.0309 | -35.0744 | -27.4768
17.037 | -35.0749 | -27.4841
17.0431 | -35.0754 | -27.4916
17.0492 | -35.0759 | -27.4989
17.0553 | -35.0765 | -27.5061
17.0614 | -35.077 | -27.5132
17.0675 | -35.0776 | -27.5205
17.0736 | -35.078 | -27.5281
17.0797 | -35.0786 | -27.5352
17.0858 | -35.0791 | -27.5428
17.0919 | -35.0796 | -27.5504
17.098 | -35.0801 | -27.558
17.1041 | -35.0806 | -27.5653
17.1102 | -35.0812 | -27.5729
17.1163 | -35.0817 | -27.5804
17.1225 | -35.0822 | -27.588
17.1286 | -35.0827 | -27.5953
17.1347 | -35.0833 | -27.603
17.1408 | -35.0838 | -27.6106
17.1469 | -35.0843 | -27.6182
17.153 | -35.0848 | -27.6255
17.1591 | -35.0853 | -27.6331
17.1652 | -35.0858 | -27.6405
17.1713 | -35.0863 | -27.6479
17.1774 | -35.0868 | -27.6553
17.1835 | -35.0874 | -27.6628
17.1896 | -35.0878 | -27.6702
17.1957 | -35.0884 | -27.6778
17.2018 | -35.0889 | -27.6851
17.2079 | -35.0894 | -27.6927
17.214 | -35.0899 | -27.7003
17.2201 | -35.0904 | -27.7077
17.2262 | -35.0909 | -27.7153
17.2323 | -35.0915 | -27.7228
17.2384 | -35.092 | -27.7301
17.2445 | -35.0925 | -27.7375
17.2506 | -35.093 | -27.745
17.2567 | -35.0935 | -27.7526
17.2628 | -35.0941 | -27.76
17.269 | -35.0946 | -27.7674
17.2751 | -35.0951 | -27.7748
17.2812 | -35.0957 | -27.7823
17.2873 | -35.0962 | -27.7897
17.2934 | -35.0967 | -27.7971
17.2995 | -35.0972 | -27.8048
17.3056 | -35.0977 | -27.8121
17.3117 | -35.0983 | -27.8196
17.3178 | -35.0988 | -27.827
17.3239 | -35.0993 | -27.8343
17.33 | -35.0998 | -27.8418
17.3361 | -35.1003 | -27.8492
17.3422 | -35.1008 | -27.8565
17.3483 | -35.1013 | -27.864
17.3544 | -35.1018 | -27.8717
17.3605 | -35.1023 | -27.8791
17.3666 | -35.1028 | -27.8865
17.3727 | -35.1034 | -27.8939
17.3788 | -35.1039 | -27.9011
17.3849 | -35.1044 | -27.9085
17.391 | -35.1049 | -27.9159
17.3971 | -35.1054 | -27.9235
17.4032 | -35.1059 | -27.931
17.4094 | -35.1065 | -27.9383
17.4155 | -35.107 | -27.9456
17.4216 | -35.1075 | -27.953
17.4277 | -35.108 | -27.9606
17.4338 | -35.1086 | -27.968
17.4399 | -35.1091 | -27.9757
17.446 | -35.1097 | -27.9829
17.4521 | -35.1102 | -27.9902
17.4582 | -35.1107 | -27.9976
17.4643 | -35.1112 | -28.0052
17.4704 | -35.1117 | -28.0126
17.4765 | -35.1123 | -28.0202
17.4826 | -35.1128 | -28.0276
17.4887 | -35.1133 | -28.0349
17.4948 | -35.1138 | -28.0423
17.5009 | -35.1143 | -28.0498
17.507 | -35.1149 | -28.0573
17.5131 | -35.1154 | -28.0649
17.5192 | -35.1159 | -28.072
17.5253 | -35.1164 | -28.0795
17.5314 | -35.1169 | -28.0869
17.5375 | -35.1174 | -28.0942
17.5436 | -35.1179 | -28.1019
17.5497 | -35.1184 | -28.1093
17.5559 | -35.1189 | -28.1166
17.562 | -35.1194 | -28.124
17.5681 | -35.12 | -28.1315
17.5742 | -35.1205 | -28.1391
17.5803 | -35.121 | -28.1464
17.5864 | -35.1215 | -28.1538
17.5925 | -35.1221 | -28.1613
17.5986 | -35.1226 | -28.1688
17.6047 | -35.1231 | -28.1764
17.6108 | -35.1236 | -28.1838
17.6169 | -35.1242 | -28.1912
17.623 | -35.1247 | -28.1985
17.6291 | -35.1252 | -28.2059
17.6352 | -35.1257 | -28.2133
17.6413 | -35.1262 | -28.2207
17.6474 | -35.1267 | -28.2282
17.6535 | -35.1272 | -28.2356
17.6596 | -35.1277 | -28.2431
17.6657 | -35.1283 | -28.2504
17.6718 | -35.1288 | -28.258
17.6779 | -35.1293 | -28.2653
17.684 | -35.1298 | -28.2727
17.6901 | -35.1303 | -28.2801
17.6963 | -35.1308 | -28.2873
17.7024 | -35.1314 | -28.2948
17.7085 | -35.1319 | -28.3021
17.7146 | -35.1324 | -28.3095
17.7207 | -35.1329 | -28.3167
17.7268 | -35.1334 | -28.3241
17.7329 | -35.1339 | -28.3315
17.739 | -35.1344 | -28.339
17.7451 | -35.135 | -28.3463
17.7512 | -35.1355 | -28.3537
17.7573 | -35.136 | -28.3613
17.7634 | -35.1365 | -28.3684
17.7695 | -35.137 | -28.3759
17.7756 | -35.1375 | -28.3834
17.7817 | -35.138 | -28.3907
17.7878 | -35.1385 | -28.3981
17.7939 | -35.139 | -28.4056
17.8 | -35.1395 | -28.4131
17.8061 | -35.1401 | -28.4206
17.8122 | -35.1406 | -28.428
17.8183 | -35.1411 | -28.4353
17.8244 | -35.1416 | -28.4428
17.8305 | -35.1421 | -28.4502
17.8366 | -35.1427 | -28.4572
17.8428 | -35.1432 | -28.4647
17.8489 | -35.1437 | -28.472
17.855 | -35.1442 | -28.4794
17.8611 | -35.1447 | -28.4867
17.8672 | -35.1452 | -28.4939
17.8733 | -35.1458 | -28.5013
17.8794 | -35.1463 | -28.5086
17.8855 | -35.1468 | -28.516
17.8916 | -35.1473 | -28.5234
17.8977 | -35.1478 | -28.531
17.9038 | -35.1483 | -28.5382
17.9099 | -35.1488 | -28.5456
17.916 | -35.1493 | -28.5531
17.9221 | -35.1498 | -28.5605
17.9282 | -35.1504 | -28.568
17.9343 | -35.1509 | -28.5755
17.9404 | -35.1513 | -28.5828
17.9465 | -35.1518 | -28.5901
17.9526 | -35.1523 | -28.5975
17.9587 | -35.1529 | -28.6048
17.9648 | -35.1534 | -28.6121
17.9709 | -35.1539 | -28.6194
17.977 | -35.1544 | -28.6268
17.9832 | -35.1549 | -28.6345
17.9893 | -35.1554 | -28.642
17.9954 | -35.1559 | -28.6494
18.0015 | -35.1564 | -28.6568
18.0076 | -35.1569 | -28.6641
18.0137 | -35.1574 | -28.6713
18.0198 | -35.1579 | -28.6787
18.0259 | -35.1584 | -28.6861
18.032 | -35.1589 | -28.6933
18.0381 | -35.1594 | -28.7007
18.0442 | -35.16 | -28.708
18.0503 | -35.1605 | -28.7155
18.0564 | -35.161 | -28.7228
18.0625 | -35.1615 | -28.7304
18.0686 | -35.162 | -28.738
18.0747 | -35.1625 | -28.7455
18.0808 | -35.163 | -28.7529
18.0869 | -35.1635 | -28.7603
18.093 | -35.164 | -28.7679
18.0991 | -35.1645 | -28.7754
18.1052 | -35.165 | -28.7828
18.1113 | -35.1656 | -28.7903
18.1174 | -35.1661 | -28.7976
18.1236 | -35.1666 | -28.8049
18.1297 | -35.1671 | -28.8123
18.1358 | -35.1676 | -28.8197
18.1419 | -35.1681 | -28.8271
18.148 | -35.1686 | -28.8344
18.1541 | -35.1692 | -28.8418
18.1602 | -35.1697 | -28.8492
18.1663 | -35.1702 | -28.8565
18.1724 | -35.1707 | -28.8639
18.1785 | -35.1712 | -28.8713
18.1846 | -35.1718 | -28.8786
18.1907 | -35.1722 | -28.8862
18.1968 | -35.1727 | -28.8936
18.2029 | -35.1733 | -28.9009
18.209 | -35.1738 | -28.9082
18.2151 | -35.1742 | -28.9157
18.2212 | -35.1748 | -28.9231
18.2273 | -35.1753 | -28.9303
18.2334 | -35.1758 | -28.9376
18.2395 | -35.1763 | -28.9451
18.2456 | -35.1769 | -28.9524
18.2517 | -35.1774 | -28.96
18.2578 | -35.1779 | -28.9672
18.2639 | -35.1784 | -28.9747
18.2701 | -35.1789 | -28.982
18.2762 | -35.1794 | -28.9892
18.2823 | -35.1799 | -28.9964
18.2884 | -35.1805 | -29.0036
18.2945 | -35.181 | -29.0107
18.3006 | -35.1815 | -29.0181
18.3067 | -35.182 | -29.0254
18.3128 | -35.1825 | -29.0327
18.3189 | -35.1829 | -29.0401
18.325 | -35.1834 | -29.0472
18.3311 | -35.1839 | -29.0545
18.3372 | -35.1844 | -29.0617
18.3433 | -35.185 | -29.0692
18.3494 | -35.1854 | -29.0765
18.3555 | -35.186 | -29.0837
18.3616 | -35.1865 | -29.091
18.3677 | -35.187 | -29.0982
18.3738 | -35.1875 | -29.1054
18.3799 | -35.188 | -29.1127
18.386 | -35.1885 | -29.1201
18.3921 | -35.189 | -29.1275
18.3982 | -35.1895 | -29.1346
18.4043 | -35.19 | -29.1419
18.4105 | -35.1905 | -29.1494
18.4166 | -35.191 | -29.1568
18.4227 | -35.1915 | -29.1642
18.4288 | -35.192 | -29.1717
18.4349 | -35.1925 | -29.1789
18.441 | -35.193 | -29.1862
18.4471 | -35.1935 | -29.1935
18.4532 | -35.194 | -29.201
18.4593 | -35.1945 | -29.2084
18.4654 | -35.195 | -29.2157
18.4715 | -35.1955 | -29.2232
18.4776 | -35.196 | -29.2306
18.4837 | -35.1965 | -29.238
18.4898 | -35.197 | -29.2455
18.4959 | -35.1975 | -29.2529
18.502 | -35.198 | -29.2603
18.5081 | -35.1985 | -29.2676
18.5142 | -35.199 | -29.275
18.5203 | -35.1995 | -29.2823
18.5264 | -35.2 | -29.2895
18.5325 | -35.2005 | -29.2968
18.5386 | -35.201 | -29.3041
18.5447 | -35.2015 | -29.3114
18.5508 | -35.202 | -29.3185
18.557 | -35.2025 | -29.3259
18.5631 | -35.203 | -29.3332
18.5692 | -35.2035 | -29.3406
18.5753 | -35.204 | -29.3477
18.5814 | -35.2045 | -29.355
18.5875 | -35.205 | -29.3626
18.5936 | -35.2055 | -29.3698
18.5997 | -35.2059 | -29.3772
18.6058 | -35.2064 | -29.3843
18.6119 | -35.2069 | -29.3916
18.618 | -35.2074 | -29.3988
18.6241 | -35.2079 | -29.406
18.6302 | -35.2084 | -29.4134
18.6363 | -35.2089 | -29.4207
18.6424 | -35.2094 | -29.4279
18.6485 | -35.2099 | -29.4351
18.6546 | -35.2104 | -29.4424
18.6607 | -35.2109 | -29.4497
18.6668 | -35.2114 | -29.4568
18.6729 | -35.2119 | -29.4642
18.679 | -35.2124 | -29.4714
18.6851 | -35.2129 | -29.4786
18.6912 | -35.2134 | -29.4859
18.6974 | -35.2139 | -29.4931
18.7035 | -35.2144 | -29.5005
18.7096 | -35.2149 | -29.5081
18.7157 | -35.2154 | -29.5155
18.7218 | -35.2159 | -29.5227
18.7279 | -35.2163 | -29.53
18.734 | -35.2168 | -29.5375
18.7401 | -35.2173 | -29.5448
18.7462 | -35.2178 | -29.5518
18.7523 | -35.2183 | -29.5593
18.7584 | -35.2188 | -29.5667
18.7645 | -35.2193 | -29.5742
18.7706 | -35.2198 | -29.5816
18.7767 | -35.2203 | -29.5889
18.7828 | -35.2207 | -29.596
18.7889 | -35.2213 | -29.6033
18.795 | -35.2218 | -29.6106
18.8011 | -35.2223 | -29.6177
18.8072 | -35.2228 | -29.6252
18.8133 | -35.2232 | -29.6326
18.8194 | -35.2237 | -29.6396
18.8255 | -35.2242 | -29.6468
18.8316 | -35.2247 | -29.654
18.8377 | -35.2252 | -29.6612
18.8439 | -35.2257 | -29.6685
18.85 | -35.2262 | -29.6759
18.8561 | -35.2266 | -29.6832
18.8622 | -35.2271 | -29.6904
18.8683 | -35.2276 | -29.6977
18.8744 | -35.2281 | -29.705
18.8805 | -35.2286 | -29.7125
18.8866 | -35.2291 | -29.72
18.8927 | -35.2296 | -29.7272
18.8988 | -35.2301 | -29.7344
18.9049 | -35.2306 | -29.7416
18.911 | -35.231 | -29.749
18.9171 | -35.2315 | -29.7563
18.9232 | -35.232 | -29.7635
18.9293 | -35.2325 | -29.771
18.9354 | -35.233 | -29.7782
18.9415 | -35.2335 | -29.7854
18.9476 | -35.234 | -29.7924
18.9537 | -35.2344 | -29.7998
18.9598 | -35.2349 | -29.8071
18.9659 | -35.2354 | -29.8144
18.972 | -35.2359 | -29.8216
18.9781 | -35.2364 | -29.8285
18.9843 | -35.2368 | -29.836
18.9904 | -35.2373 | -29.8433
18.9965 | -35.2378 | -29.8505
19.0026 | -35.2383 | -29.8577
19.0087 | -35.2388 | -29.865
19.0148 | -35.2393 | -29.8726
19.0209 | -35.2398 | -29.8798
19.027 | -35.2403 | -29.8874
19.0331 | -35.2407 | -29.8947
19.0392 | -35.2413 | -29.9019
19.0453 | -35.2417 | -29.9093
19.0514 | -35.2422 | -29.9166
19.0575 | -35.2427 | -29.9239
19.0636 | -35.2432 | -29.9309
19.0697 | -35.2437 | -29.9381
19.0758 | -35.2442 | -29.9454
19.0819 | -35.2448 | -29.9526
19.088 | -35.2452 | -29.96
19.0941 | -35.2457 | -29.9673
19.1002 | -35.2461 | -29.9744
19.1063 | -35.2466 | -29.9817
19.1124 | -35.2471 | -29.989
19.1185 | -35.2475 | -29.9963
19.1246 | -35.248 | -30.0035
19.1308 | -35.2485 | -30.0109
19.1369 | -35.249 | -30.018
19.143 | -35.2495 | -30.0253
19.1491 | -35.2499 | -30.0328
19.1552 | -35.2504 | -30.0401
19.1613 | -35.2509 | -30.0475
19.1674 | -35.2514 | -30.0548
19.1735 | -35.2519 | -30.0618
19.1796 | -35.2524 | -30.0689
19.1857 | -35.2528 | -30.0762
19.1918 | -35.2533 | -30.0836
19.1979 | -35.2538 | -30.0908
19.204 | -35.2543 | -30.0978
19.2101 | -35.2548 | -30.105
19.2162 | -35.2553 | -30.1121
19.2223 | -35.2558 | -30.1195
19.2284 | -35.2562 | -30.1268
19.2345 | -35.2567 | -30.1341
19.2406 | -35.2572 | -30.1411
19.2467 | -35.2577 | -30.1484
19.2528 | -35.2582 | -30.1557
19.2589 | -35.2587 | -30.1628
19.265 | -35.2592 | -30.1701
19.2712 | -35.2597 | -30.1773
19.2773 | -35.2602 | -30.1845
19.2834 | -35.2606 | -30.1916
19.2895 | -35.2612 | -30.1988
19.2956 | -35.2616 | -30.2063
19.3017 | -35.2621 | -30.2134
19.3078 | -35.2626 | -30.2207
19.3139 | -35.2631 | -30.2276
19.32 | -35.2636 | -30.2347
19.3261 | -35.264 | -30.2418
19.3322 | -35.2645 | -30.2488
19.3383 | -35.265 | -30.256
19.3444 | -35.2655 | -30.2631
19.3505 | -35.266 | -30.2702
19.3566 | -35.2664 | -30.2772
19.3627 | -35.2669 | -30.2844
19.3688 | -35.2674 | -30.2915
19.3749 | -35.2679 | -30.2987
19.381 | -35.2684 | -30.3057
19.3871 | -35.2689 | -30.3128
19.3932 | -35.2694 | -30.32
19.3993 | -35.2698 | -30.3273
19.4054 | -35.2703 | -30.3346
19.4115 | -35.2708 | -30.3419
19.4177 | -35.2713 | -30.3488
19.4238 | -35.2718 | -30.356
19.4299 | -35.2723 | -30.3633
19.436 | -35.2728 | -30.3708
19.4421 | -35.2732 | -30.3779
19.4482 | -35.2737 | -30.3851
19.4543 | -35.2742 | -30.3922
19.4604 | -35.2747 | -30.3995
19.4665 | -35.2752 | -30.4065
19.4726 | -35.2757 | -30.4136
19.4787 | -35.2761 | -30.4209
19.4848 | -35.2766 | -30.4282
19.4909 | -35.2771 | -30.4352
19.497 | -35.2776 | -30.4424
19.5031 | -35.2781 | -30.4494
19.5092 | -35.2785 | -30.4566
19.5153 | -35.279 | -30.4639
19.5214 | -35.2795 | -30.4711
19.5275 | -35.2799 | -30.4783
19.5336 | -35.2804 | -30.4854
19.5397 | -35.2809 | -30.4927
19.5458 | -35.2814 | -30.5001
19.5519 | -35.2819 | -30.5072
19.5581 | -35.2823 | -30.5141
19.5642 | -35.2828 | -30.5213
19.5703 | -35.2832 | -30.5285
19.5764 | -35.2838 | -30.5358
19.5825 | -35.2842 | -30.5431
19.5886 | -35.2847 | -30.5501
19.5947 | -35.2851 | -30.557
19.6008 | -35.2856 | -30.5642
19.6069 | -35.2861 | -30.5713
19.613 | -35.2866 | -30.5785
19.6191 | -35.2871 | -30.5858
19.6252 | -35.2876 | -30.5931
19.6313 | -35.2881 | -30.6
19.6374 | -35.2885 | -30.6071
19.6435 | -35.289 | -30.6141
19.6496 | -35.2895 | -30.6212
19.6557 | -35.2899 | -30.6283
19.6618 | -35.2904 | -30.6353
19.6679 | -35.2908 | -30.6425
19.674 | -35.2913 | -30.6497
19.6801 | -35.2918 | -30.6568
19.6862 | -35.2923 | -30.664
19.6923 | -35.2928 | -30.6709
19.6984 | -35.2933 | -30.6779
19.7046 | -35.2938 | -30.6851
19.7107 | -35.2942 | -30.6923
19.7168 | -35.2947 | -30.6996
19.7229 | -35.2951 | -30.7068
19.729 | -35.2956 | -30.7141
19.7351 | -35.296 | -30.7212
19.7412 | -35.2965 | -30.7283
19.7473 | -35.297 | -30.7355
19.7534 | -35.2975 | -30.7427
19.7595 | -35.2979 | -30.7499
19.7656 | -35.2984 | -30.7572
19.7717 | -35.2988 | -30.7644
19.7778 | -35.2993 | -30.7715
19.7839 | -35.2998 | -30.7785
19.79 | -35.3003 | -30.7856
19.7961 | -35.3007 | -30.7927
19.8022 | -35.3012 | -30.7999
19.8083 | -35.3017 | -30.8071
19.8144 | -35.3022 | -30.8143
19.8205 | -35.3027 | -30.8212
19.8266 | -35.3032 | -30.8283
19.8327 | -35.3036 | -30.8354
19.8388 | -35.3041 | -30.8427
19.845 | -35.3045 | -30.8498
19.8511 | -35.305 | -30.8568
19.8572 | -35.3055 | -30.8637
19.8633 | -35.3059 | -30.8707
19.8694 | -35.3064 | -30.8779
19.8755 | -35.3068 | -30.8849
19.8816 | -35.3073 | -30.892
19.8877 | -35.3077 | -30.8994
19.8938 | -35.3082 | -30.9063
19.8999 | -35.3087 | -30.9136
19.906 | -35.3091 | -30.9209
19.9121 | -35.3097 | -30.9275
19.9182 | -35.3101 | -30.9345
19.9243 | -35.3106 | -30.9415
19.9304 | -35.3111 | -30.9485
19.9365 | -35.3115 | -30.9557
19.9426 | -35.312 | -30.9622
19.9487 | -35.3125 | -30.9696
19.9548 | -35.313 | -30.9767
19.9609 | -35.3134 | -30.9837
19.967 | -35.314 | -30.9909
19.9731 | -35.3144 | -30.998
19.9792 | -35.3148 | -31.005
19.9853 | -35.3153 | -31.0121
19.9915 | -35.3158 | -31.0193
19.9976 | -35.3163 | -31.0264
20.0037 | -35.3168 | -31.0335
20.0098 | -35.3173 | -31.0407
20.0159 | -35.3177 | -31.0478
20.022 | -35.3182 | -31.055
20.0281 | -35.3186 | -31.0621
20.0342 | -35.3191 | -31.0693
20.0403 | -35.3196 | -31.0764
20.0464 | -35.32 | -31.0835
20.0525 | -35.3205 | -31.0904
20.0586 | -35.3209 | -31.0976
20.0647 | -35.3214 | -31.1047
20.0708 | -35.3218 | -31.1118
20.0769 | -35.3223 | -31.1189
20.083 | -35.3228 | -31.1258
20.0891 | -35.3233 | -31.1328
20.0952 | -35.3238 | -31.1397
20.1013 | -35.3242 | -31.1468
20.1074 | -35.3247 | -31.1539
20.1135 | -35.3252 | -31.161
20.1196 | -35.3257 | -31.1676
20.1257 | -35.3261 | -31.1747
20.1319 | -35.3266 | -31.1817
20.138 | -35.3271 | -31.1886
20.1441 | -35.3275 | -31.1956
20.1502 | -35.328 | -31.2028
20.1563 | -35.3285 | -31.2096
20.1624 | -35.329 | -31.2164
20.1685 | -35.3294 | -31.2234
20.1746 | -35.3299 | -31.2303
20.1807 | -35.3303 | -31.2374
20.1868 | -35.3308 | -31.2445
20.1929 | -35.3312 | -31.2516
20.199 | -35.3317 | -31.2587
20.2051 | -35.3321 | -31.2656
20.2112 | -35.3326 | -31.2725
20.2173 | -35.3331 | -31.2793
20.2234 | -35.3335 | -31.2864
20.2295 | -35.334 | -31.2931
20.2356 | -35.3345 | -31.3002
20.2417 | -35.335 | -31.3073
20.2478 | -35.3354 | -31.3143
20.2539 | -35.3359 | -31.3215
20.26 | -35.3363 | -31.3286
20.2661 | -35.3368 | -31.3357
20.2723 | -35.3373 | -31.3427
20.2784 | -35.3377 | -31.3497
20.2845 | -35.3382 | -31.3567
20.2906 | -35.3387 | -31.3638
20.2967 | -35.3391 | -31.3706
20.3028 | -35.3396 | -31.3774
20.3089 | -35.34 | -31.3844
20.315 | -35.3405 | -31.3912
20.3211 | -35.3409 | -31.3983
20.3272 | -35.3414 | -31.4054
20.3333 | -35.3419 | -31.4123
20.3394 | -35.3424 | -31.4193
20.3455 | -35.3429 | -31.4262
20.3516 | -35.3434 | -31.4332
20.3577 | -35.3438 | -31.4401
20.3638 | -35.3442 | -31.4471
20.3699 | -35.3447 | -31.4538
20.376 | -35.3452 | -31.4605
20.3821 | -35.3457 | -31.4674
20.3882 | -35.3461 | -31.4744
20.3943 | -35.3466 | -31.4814
20.4004 | -35.347 | -31.4881
20.4065 | -35.3475 | -31.4952
20.4126 | -35.348 | -31.5019
20.4188 | -35.3484 | -31.5088
20.4249 | -35.3489 | -31.5156
20.431 | -35.3494 | -31.5223
20.4371 | -35.3499 | -31.5291
20.4432 | -35.3503 | -31.5361
20.4493 | -35.3507 | -31.5432
20.4554 | -35.3512 | -31.55
20.4615 | -35.3517 | -31.5569
20.4676 | -35.3521 | -31.564
20.4737 | -35.3526 | -31.5709
20.4798 | -35.353 | -31.5779
20.4859 | -35.3535 | -31.5852
20.492 | -35.354 | -31.5919
20.4981 | -35.3544 | -31.599
20.5042 | -35.3549 | -31.606
20.5103 | -35.3553 | -31.613
20.5164 | -35.3557 | -31.6197
20.5225 | -35.3562 | -31.6265
20.5286 | -35.3566 | -31.6332
20.5347 | -35.3571 | -31.6403
20.5408 | -35.3575 | -31.6468
20.5469 | -35.358 | -31.6541
20.553 | -35.3584 | -31.6613
20.5592 | -35.3589 | -31.668
20.5653 | -35.3594 | -31.675
20.5714 | -35.3598 | -31.6822
20.5775 | -35.3603 | -31.6893
20.5836 | -35.3608 | -31.6963
20.5897 | -35.3612 | -31.7033
20.5958 | -35.3616 | -31.7103
20.6019 | -35.3621 | -31.717
20.608 | -35.3625 | -31.7241
20.6141 | -35.363 | -31.7308
20.6202 | -35.3634 | -31.7376
20.6263 | -35.3639 | -31.7443
20.6324 | -35.3643 | -31.751
20.6385 | -35.3648 | -31.758
20.6446 | -35.3653 | -31.7652
20.6507 | -35.3658 | -31.7719
20.6568 | -35.3662 | -31.7788
20.6629 | -35.3666 | -31.7859
20.669 | -35.367 | -31.7929
20.6751 | -35.3675 | -31.7998
20.6812 | -35.3679 | -31.8064
20.6873 | -35.3684 | -31.8134
20.6934 | -35.3688 | -31.8202
20.6995 | -35.3693 | -31.8271
20.7057 | -35.3698 | -31.8339
20.7118 | -35.3702 | -31.8408
20.7179 | -35.3707 | -31.8475
20.724 | -35.3711 | -31.8545
20.7301 | -35.3716 | -31.861
20.7362 | -35.3721 | -31.8679
20.7423 | -35.3726 | -31.875
20.7484 | -35.3731 | -31.882
20.7545 | -35.3736 | -31.8886
20.7606 | -35.374 | -31.8957
20.7667 | -35.3745 | -31.9027
20.7728 | -35.375 | -31.9096
20.7789 | -35.3754 | -31.9163
20.785 | -35.3759 | -31.9231
20.7911 | -35.3764 | -31.9299
20.7972 | -35.3768 | -31.9365
20.8033 | -35.3773 | -31.9436
20.8094 | -35.3778 | -31.9507
20.8155 | -35.3783 | -31.9574
20.8216 | -35.3787 | -31.9643
20.8277 | -35.3791 | -31.9711
20.8338 | -35.3796 | -31.9781
20.8399 | -35.3801 | -31.9853
20.8461 | -35.3805 | -31.992
20.8522 | -35.3811 | -31.9989
20.8583 | -35.3815 | -32.0056
20.8644 | -35.382 | -32.0126
20.8705 | -35.3824 | -32.0196
20.8766 | -35.3828 | -32.0267
20.8827 | -35.3833 | -32.0337
20.8888 | -35.3838 | -32.0407
20.8949 | -35.3842 | -32.0471
20.901 | -35.3846 | -32.0535
20.9071 | -35.3851 | -32.0602
20.9132 | -35.3856 | -32.0669
20.9193 | -35.386 | -32.0741
20.9254 | -35.3865 | -32.0807
20.9315 | -35.3869 | -32.0872
20.9376 | -35.3873 | -32.094
20.9437 | -35.3877 | -32.1008
20.9498 | -35.3882 | -32.1077
20.9559 | -35.3886 | -32.1147
20.962 | -35.3891 | -32.1214
20.9681 | -35.3896 | -32.128
20.9742 | -35.3901 | -32.1348
20.9803 | -35.3905 | -32.1417
20.9864 | -35.3909 | -32.1483
20.9926 | -35.3915 | -32.1549
20.9987 | -35.3919 | -32.1619
21.0048 | -35.3923 | -32.169
21.0109 | -35.3929 | -32.176
21.017 | -35.3933 | -32.183
21.0231 | -35.3938 | -32.1898
21.0292 | -35.3942 | -32.1965
21.0353 | -35.3946 | -32.203
21.0414 | -35.3951 | -32.2094
21.0475 | -35.3956 | -32.216
21.0536 | -35.3961 | -32.2228
21.0597 | -35.3965 | -32.2296
21.0658 | -35.3969 | -32.2361
21.0719 | -35.3974 | -32.2427
21.078 | -35.3979 | -32.2496
21.0841 | -35.3984 | -32.2563
21.0902 | -35.3989 | -32.263
21.0963 | -35.3993 | -32.2693
21.1024 | -35.3998 | -32.2767
21.1085 | -35.4002 | -32.2827
21.1146 | -35.4006 | -32.2897
21.1207 | -35.401 | -32.2964
21.1268 | -35.4015 | -32.3032
21.133 | -35.4019 | -32.3099
21.1391 | -35.4024 | -32.3164
21.1452 | -35.4028 | -32.3237
21.1513 | -35.4033 | -32.3302
21.1574 | -35.4037 | -32.337
21.1635 | -35.4042 | -32.344
21.1696 | -35.4046 | -32.3508
21.1757 | -35.4051 | -32.3571
21.1818 | -35.4055 | -32.3638
21.1879 | -35.406 | -32.3705
21.194 | -35.4064 | -32.3772
21.2001 | -35.4069 | -32.384
21.2062 | -35.4074 | -32.3908
21.2123 | -35.4078 | -32.3971
21.2184 | -35.4083 | -32.4042
21.2245 | -35.4087 | -32.4107
21.2306 | -35.4092 | -32.4176
21.2367 | -35.4096 | -32.4245
21.2428 | -35.4102 | -32.4313
21.2489 | -35.4106 | -32.4379
21.255 | -35.4111 | -32.445
21.2611 | -35.4116 | -32.4512
21.2672 | -35.4121 | -32.458
21.2733 | -35.4125 | -32.4645
21.2795 | -35.4129 | -32.4708
21.2856 | -35.4134 | -32.4772
21.2917 | -35.4139 | -32.4843
21.2978 | -35.4143 | -32.4912
21.3039 | -35.4148 | -32.4981
21.31 | -35.4153 | -32.5052
21.3161 | -35.4157 | -32.5118
21.3222 | -35.4161 | -32.5185
21.3283 | -35.4167 | -32.5252
21.3344 | -35.4171 | -32.5319
21.3405 | -35.4175 | -32.5392
21.3466 | -35.418 | -32.5459
21.3527 | -35.4184 | -32.5526
21.3588 | -35.4189 | -32.5592
21.3649 | -35.4194 | -32.5657
21.371 | -35.4198 | -32.5724
21.3771 | -35.4203 | -32.5791
21.3832 | -35.4208 | -32.5858
21.3893 | -35.4212 | -32.5921
21.3954 | -35.4217 | -32.599
21.4015 | -35.4221 | -32.6054
21.4076 | -35.4225 | -32.6122
21.4137 | -35.423 | -32.6187
21.4199 | -35.4234 | -32.6255
21.426 | -35.4239 | -32.6319
21.4321 | -35.4243 | -32.6388
21.4382 | -35.4247 | -32.6453
21.4443 | -35.4252 | -32.6518
21.4504 | -35.4256 | -32.6584
21.4565 | -35.4261 | -32.6648
21.4626 | -35.4265 | -32.6716
21.4687 | -35.427 | -32.6782
21.4748 | -35.4275 | -32.6849
21.4809 | -35.428 | -32.6914
21.487 | -35.4285 | -32.6982
21.4931 | -35.4289 | -32.7044
21.4992 | -35.4294 | -32.7113
21.5053 | -35.4298 | -32.718
21.5114 | -35.4303 | -32.7244
21.5175 | -35.4308 | -32.7312
21.5236 | -35.4312 | -32.7376
21.5297 | -35.4317 | -32.7441
21.5358 | -35.4321 | -32.7507
21.5419 | -35.4325 | -32.7576
21.548 | -35.433 | -32.7641
21.5541 | -35.4335 | -32.7709
21.5602 | -35.4339 | -32.7777
21.5664 | -35.4344 | -32.7845
21.5725 | -35.4349 | -32.7913
21.5786 | -35.4353 | -32.7983
21.5847 | -35.4358 | -32.8049
21.5908 | -35.4361 | -32.8117
21.5969 | -35.4366 | -32.8182
21.603 | -35.437 | -32.8251
21.6091 | -35.4375 | -32.8319
21.6152 | -35.4379 | -32.8384
21.6213 | -35.4384 | -32.8451
21.6274 | -35.4388 | -32.8516
21.6335 | -35.4393 | -32.8578
21.6396 | -35.4397 | -32.8645
21.6457 | -35.4402 | -32.8711
21.6518 | -35.4406 | -32.8773
21.6579 | -35.441 | -32.8839
21.664 | -35.4415 | -32.8904
21.6701 | -35.442 | -32.8965
21.6762 | -35.4424 | -32.9032
21.6823 | -35.4429 | -32.9097
21.6884 | -35.4434 | -32.9168
21.6945 | -35.4438 | -32.9234
21.7006 | -35.4443 | -32.9302
21.7068 | -35.4446 | -32.9366
21.7129 | -35.4451 | -32.943
21.719 | -35.4455 | -32.9498
21.7251 | -35.4459 | -32.9562
21.7312 | -35.4464 | -32.9628
21.7373 | -35.4469 | -32.9685
21.7434 | -35.4473 | -32.9752
21.7495 | -35.4478 | -32.9815
21.7556 | -35.4482 | -32.988
21.7617 | -35.4487 | -32.9945
21.7678 | -35.4492 | -33.0006
21.7739 | -35.4496 | -33.0067
21.78 | -35.45 | -33.0133
21.7861 | -35.4504 | -33.0201
21.7922 | -35.4509 | -33.0266
21.7983 | -35.4513 | -33.0328
21.8044 | -35.4518 | -33.0392
21.8105 | -35.4522 | -33.045
21.8166 | -35.4526 | -33.0515
21.8227 | -35.4531 | -33.0581
21.8288 | -35.4535 | -33.0645
21.8349 | -35.4539 | -33.0714
21.841 | -35.4544 | -33.0781
21.8471 | -35.4548 | -33.0842
21.8533 | -35.4553 | -33.0909
21.8594 | -35.4557 | -33.0974
21.8655 | -35.4562 | -33.104
21.8716 | -35.4566 | -33.1106
21.8777 | -35.4571 | -33.1168
21.8838 | -35.4576 | -33.1234
21.8899 | -35.458 | -33.1299
21.896 | -35.4585 | -33.136
21.9021 | -35.4589 | -33.1428
21.9082 | -35.4594 | -33.1495
21.9143 | -35.4598 | -33.1558
21.9204 | -35.4603 | -33.1624
21.9265 | -35.4608 | -33.1689
21.9326 | -35.4612 | -33.1755
21.9387 | -35.4617 | -33.1815
21.9448 | -35.4621 | -33.1883
21.9509 | -35.4626 | -33.1946
21.957 | -35.463 | -33.201
21.9631 | -35.4635 | -33.2079
21.9692 | -35.464 | -33.2145
21.9753 | -35.4645 | -33.221
21.9814 | -35.4649 | -33.2273
21.9875 | -35.4654 | -33.2334
21.9937 | -35.4659 | -33.2397
21.9998 | -35.4664 | -33.2461
22.0059 | -35.4668 | -33.2526
22.012 | -35.4673 | -33.259
22.0181 | -35.4678 | -33.2653
22.0242 | -35.4682 | -33.2715
22.0303 | -35.4686 | -33.2783
22.0364 | -35.4691 | -33.2848
22.0425 | -35.4696 | -33.2916
22.0486 | -35.47 | -33.2982
22.0547 | -35.4704 | -33.3045
22.0608 | -35.4709 | -33.311
22.0669 | -35.4713 | -33.3173
22.073 | -35.4718 | -33.3239
22.0791 | -35.4722 | -33.3306
22.0852 | -35.4727 | -33.3366
22.0913 | -35.4732 | -33.3429
22.0974 | -35.4736 | -33.3495
22.1035 | -35.474 | -33.356
22.1096 | -35.4744 | -33.3623
22.1157 | -35.4749 | -33.3689
22.1218 | -35.4753 | -33.3752
22.1279 | -35.4757 | -33.3808
22.134 | -35.4762 | -33.3875
22.1402 | -35.4767 | -33.3938
22.1463 | -35.4772 | -33.3999
22.1524 | -35.4777 | -33.4065
22.1585 | -35.4781 | -33.4131
22.1646 | -35.4786 | -33.4188
22.1707 | -35.479 | -33.4254
22.1768 | -35.4795 | -33.4318
22.1829 | -35.4799 | -33.4381
22.189 | -35.4805 | -33.4447
22.1951 | -35.4808 | -33.4508
22.2012 | -35.4812 | -33.457
22.2073 | -35.4817 | -33.4635
22.2134 | -35.4821 | -33.47
22.2195 | -35.4825 | -33.4761
22.2256 | -35.483 | -33.4828
22.2317 | -35.4834 | -33.4889
22.2378 | -35.4838 | -33.4954
22.2439 | -35.4843 | -33.5018
22.25 | -35.4848 | -33.508
22.2561 | -35.4852 | -33.514
22.2622 | -35.4857 | -33.5203
22.2683 | -35.4862 | -33.5266
22.2744 | -35.4866 | -33.5331
22.2806 | -35.4871 | -33.5395
22.2867 | -35.4876 | -33.5457
22.2928 | -35.488 | -33.552
22.2989 | -35.4884 | -33.5586
22.305 | -35.4889 | -33.5644
22.3111 | -35.4894 | -33.5709
22.3172 | -35.4898 | -33.5773
22.3233 | -35.4903 | -33.584
22.3294 | -35.4908 | -33.5903
22.3355 | -35.4912 | -33.5968
22.3416 | -35.4917 | -33.6031
22.3477 | -35.4921 | -33.6095
22.3538 | -35.4926 | -33.6159
22.3599 | -35.4931 | -33.6224
22.366 | -35.4935 | -33.6288
22.3721 | -35.494 | -33.6353
22.3782 | -35.4945 | -33.6414
22.3843 | -35.495 | -33.6478
22.3904 | -35.4954 | -33.6541
22.3965 | -35.4958 | -33.6604
22.4026 | -35.4963 | -33.6668
22.4087 | -35.4968 | -33.6727
22.4148 | -35.4972 | -33.679
22.4209 | -35.4976 | -33.6854
22.4271 | -35.4981 | -33.6919
22.4332 | -35.4985 | -33.6985
22.4393 | -35.499 | -33.7048
22.4454 | -35.4994 | -33.7109
22.4515 | -35.4999 | -33.7173
22.4576 | -35.5004 | -33.7239
22.4637 | -35.5009 | -33.7301
22.4698 | -35.5013 | -33.7365
22.4759 | -35.5018 | -33.7427
22.482 | -35.5022 | -33.749
22.4881 | -35.5027 | -33.7553
22.4942 | -35.5032 | -33.7616
22.5003 | -35.5036 | -33.768
22.5064 | -35.5041 | -33.774
22.5125 | -35.5046 | -33.7797
22.5186 | -35.505 | -33.7863
22.5247 | -35.5054 | -33.7927
22.5308 | -35.5059 | -33.799
22.5369 | -35.5063 | -33.8057
22.543 | -35.5067 | -33.812
22.5491 | -35.5072 | -33.818
22.5552 | -35.5077 | -33.8245
22.5613 | -35.5082 | -33.8309
22.5675 | -35.5087 | -33.8375
22.5736 | -35.5091 | -33.8438
22.5797 | -35.5095 | -33.8502
22.5858 | -35.51 | -33.8564
22.5919 | -35.5104 | -33.8623
22.598 | -35.5108 | -33.869
22.6041 | -35.5113 | -33.8753
22.6102 | -35.5118 | -33.882
22.6163 | -35.5122 | -33.8879
22.6224 | -35.5127 | -33.8941
22.6285 | -35.5132 | -33.9003
22.6346 | -35.5136 | -33.9068
22.6407 | -35.514 | -33.9131
22.6468 | -35.5145 | -33.9192
22.6529 | -35.5149 | -33.9251
22.659 | -35.5154 | -33.9316
22.6651 | -35.5158 | -33.938
22.6712 | -35.5163 | -33.9443
22.6773 | -35.5167 | -33.9506
22.6834 | -35.5172 | -33.9569
22.6895 | -35.5177 | -33.9628
22.6956 | -35.5181 | -33.9687
22.7017 | -35.5185 | -33.975
22.7079 | -35.519 | -33.9813
22.714 | -35.5195 | -33.9875
22.7201 | -35.5199 | -33.9939
22.7262 | -35.5204 | -33.9998
22.7323 | -35.5208 | -34.0058
22.7384 | -35.5213 | -34.012
22.7445 | -35.5217 | -34.0183
22.7506 | -35.5222 | -34.0246
22.7567 | -35.5226 | -34.0307
22.7628 | -35.5231 | -34.0366
22.7689 | -35.5236 | -34.0425
22.775 | -35.524 | -34.0488
22.7811 | -35.5245 | -34.055
22.7872 | -35.525 | -34.0613
22.7933 | -35.5254 | -34.0674
22.7994 | -35.5259 | -34.0736
22.8055 | -35.5264 | -34.0798
22.8116 | -35.5268 | -34.0859
22.8177 | -35.5273 | -34.0923
22.8238 | -35.5277 | -34.0986
22.8299 | -35.5282 | -34.1045
22.836 | -35.5286 | -34.1105
22.8421 | -35.5291 | -34.1168
22.8482 | -35.5296 | -34.123
22.8544 | -35.53 | -34.1292
22.8605 | -35.5305 | -34.1354
22.8666 | -35.5309 | -34.1415
22.8727 | -35.5313 | -34.1476
22.8788 | -35.5318 | -34.1538
22.8849 | -35.5322 | -34.1598
22.891 | -35.5326 | -34.1658
22.8971 | -35.5331 | -34.1719
22.9032 | -35.5336 | -34.1779
22.9093 | -35.534 | -34.1841
22.9154 | -35.5345 | -34.1903
22.9215 | -35.5349 | -34.1967
22.9276 | -35.5353 | -34.2032
22.9337 | -35.5358 | -34.2089
22.9398 | -35.5362 | -34.2151
22.9459 | -35.5367 | -34.2213
22.952 | -35.5371 | -34.2274
22.9581 | -35.5376 | -34.2336
22.9642 | -35.538 | -34.2397
22.9703 | -35.5385 | -34.2457
22.9764 | -35.5389 | -34.2518
22.9825 | -35.5394 | -34.2578
22.9886 | -35.5398 | -34.264
22.9948 | -35.5402 | -34.2702
23.0009 | -35.5407 | -34.2765
23.007 | -35.5411 | -34.2824
23.0131 | -35.5416 | -34.2886
23.0192 | -35.542 | -34.2947
23.0253 | -35.5424 | -34.301
23.0314 | -35.5429 | -34.3071
23.0375 | -35.5433 | -34.3135
23.0436 | -35.5438 | -34.3195
23.0497 | -35.5442 | -34.3258
23.0558 | -35.5447 | -34.3322
23.0619 | -35.5451 | -34.3382
23.068 | -35.5456 | -34.3445
23.0741 | -35.546 | -34.3504
23.0802 | -35.5464 | -34.3561
23.0863 | -35.5469 | -34.3622
23.0924 | -35.5473 | -34.3682
23.0985 | -35.5478 | -34.3741
23.1046 | -35.5482 | -34.3803
23.1107 | -35.5487 | -34.3865
23.1168 | -35.5491 | -34.3929
23.1229 | -35.5495 | -34.3991
23.129 | -35.55 | -34.4053
23.1351 | -35.5505 | -34.4114
23.1413 | -35.5509 | -34.4175
23.1474 | -35.5513 | -34.4234
23.1535 | -35.5518 | -34.4294
23.1596 | -35.5522 | -34.4356
23.1657 | -35.5527 | -34.4417
23.1718 | -35.5531 | -34.4478
23.1779 | -35.5536 | -34.4542
23.184 | -35.554 | -34.4599
23.1901 | -35.5545 | -34.4661
23.1962 | -35.555 | -34.4722
23.2023 | -35.5554 | -34.4785
23.2084 | -35.5559 | -34.4847
23.2145 | -35.5563 | -34.4906
23.2206 | -35.5567 | -34.4967
23.2267 | -35.5572 | -34.5026
23.2328 | -35.5576 | -34.5086
23.2389 | -35.558 | -34.5146
23.245 | -35.5585 | -34.5208
23.2511 | -35.559 | -34.5265
23.2572 | -35.5594 | -34.5325
23.2633 | -35.5599 | -34.5386
23.2694 | -35.5603 | -34.5447
23.2755 | -35.5607 | -34.5508
23.2817 | -35.5612 | -34.5568
23.2878 | -35.5616 | -34.5627
23.2939 | -35.562 | -34.5688
23.3 | -35.5625 | -34.5749
23.3061 | -35.5629 | -34.581
23.3122 | -35.5633 | -34.5872
23.3183 | -35.5638 | -34.5932
23.3244 | -35.5642 | -34.5993
23.3305 | -35.5647 | -34.6053
23.3366 | -35.5651 | -34.6114
23.3427 | -35.5655 | -34.6175
23.3488 | -35.566 | -34.6237
23.3549 | -35.5665 | -34.6299
23.361 | -35.567 | -34.6357
23.3671 | -35.5674 | -34.6419
23.3732 | -35.5679 | -34.6481
23.3793 | -35.5683 | -34.6541
23.3854 | -35.5688 | -34.6603
23.3915 | -35.5693 | -34.6661
23.3976 | -35.5697 | -34.6721
23.4037 | -35.5701 | -34.6783
23.4098 | -35.5706 | -34.6843
23.4159 | -35.571 | -34.6906
23.422 | -35.5714 | -34.6965
23.4282 | -35.5719 | -34.7021
23.4343 | -35.5723 | -34.7081
23.4404 | -35.5728 | -34.7144
23.4465 | -35.5732 | -34.7205
23.4526 | -35.5737 | -34.7267
23.4587 | -35.5741 | -34.7328
23.4648 | -35.5746 | -34.7384
23.4709 | -35.5751 | -34.7446
23.477 | -35.5755 | -34.7508
23.4831 | -35.5759 | -34.7569
23.4892 | -35.5764 | -34.7628
23.4953 | -35.5769 | -34.7689
23.5014 | -35.5773 | -34.7749
23.5075 | -35.5777 | -34.781
23.5136 | -35.5782 | -34.787
23.5197 | -35.5787 | -34.7929
23.5258 | -35.5791 | -34.7989
23.5319 | -35.5796 | -34.8046
23.538 | -35.58 | -34.8108
23.5441 | -35.5804 | -34.8168
23.5502 | -35.5809 | -34.823
23.5563 | -35.5813 | -34.829
23.5624 | -35.5817 | -34.8351
23.5686 | -35.5822 | -34.8409
23.5747 | -35.5826 | -34.847
23.5808 | -35.5831 | -34.8529
23.5869 | -35.5836 | -34.8589
23.593 | -35.584 | -34.8649
23.5991 | -35.5844 | -34.8708
23.6052 | -35.5849 | -34.8768
23.6113 | -35.5853 | -34.8831
23.6174 | -35.5857 | -34.8892
23.6235 | -35.5862 | -34.895
23.6296 | -35.5866 | -34.9011
23.6357 | -35.5871 | -34.9073
23.6418 | -35.5875 | -34.9133
23.6479 | -35.588 | -34.9191
23.654 | -35.5884 | -34.9252
23.6601 | -35.5889 | -34.9312
23.6662 | -35.5893 | -34.9372
23.6723 | -35.5897 | -34.9429
23.6784 | -35.5902 | -34.949
23.6845 | -35.5906 | -34.9549
23.6906 | -35.591 | -34.9609
23.6967 | -35.5915 | -34.967
23.7028 | -35.5919 | -34.9729
23.7089 | -35.5924 | -34.9785
23.7151 | -35.5928 | -34.9844
23.7212 | -35.5933 | -34.9905
23.7273 | -35.5937 | -34.9967
23.7334 | -35.5942 | -35.0025
23.7395 | -35.5946 | -35.0086
23.7456 | -35.595 | -35.0144
23.7517 | -35.5955 | -35.0203
23.7578 | -35.5959 | -35.0262
23.7639 | -35.5964 | -35.0323
23.77 | -35.5968 | -35.0383
23.7761 | -35.5973 | -35.0443
23.7822 | -35.5977 | -35.0501
23.7883 | -35.5982 | -35.0561
23.7944 | -35.5986 | -35.0619
23.8005 | -35.599 | -35.0681
23.8066 | -35.5995 | -35.074
23.8127 | -35.5999 | -35.0799
23.8188 | -35.6004 | -35.086
23.8249 | -35.6008 | -35.0919
23.831 | -35.6013 | -35.0979
23.8371 | -35.6017 | -35.1039
23.8432 | -35.6021 | -35.1101
23.8493 | -35.6026 | -35.1158
23.8555 | -35.603 | -35.1215
23.8616 | -35.6035 | -35.1276
23.8677 | -35.6039 | -35.1335
23.8738 | -35.6043 | -35.1393
23.8799 | -35.6048 | -35.1453
23.886 | -35.6052 | -35.151
23.8921 | -35.6056 | -35.1571
23.8982 | -35.606 | -35.163
23.9043 | -35.6064 | -35.169
23.9104 | -35.6069 | -35.1751
23.9165 | -35.6073 | -35.181
23.9226 | -35.6077 | -35.1867
23.9287 | -35.6082 | -35.1927
23.9348 | -35.6086 | -35.1988
23.9409 | -35.609 | -35.2047
23.947 | -35.6095 | -35.2106
23.9531 | -35.6099 | -35.2163
23.9592 | -35.6103 | -35.2221
23.9653 | -35.6108 | -35.2282
23.9714 | -35.6112 | -35.2341
23.9775 | -35.6116 | -35.24
23.9836 | -35.6121 | -35.246
23.9897 | -35.6126 | -35.252
23.9958 | -35.613 | -35.2579
24.002 | -35.6134 | -35.2638
24.0081 | -35.6138 | -35.2697
24.0142 | -35.6143 | -35.2757
24.0203 | -35.6147 | -35.2816
24.0264 | -35.6152 | -35.2875
24.0325 | -35.6156 | -35.2935
24.0386 | -35.616 | -35.2996
24.0447 | -35.6164 | -35.3056
24.0508 | -35.6169 | -35.3114
24.0569 | -35.6173 | -35.3174
24.063 | -35.6178 | -35.323
24.0691 | -35.6182 | -35.329
24.0752 | -35.6187 | -35.3349
24.0813 | -35.6191 | -35.3407
24.0874 | -35.6195 | -35.3465
24.0935 | -35.6199 | -35.3526
24.0996 | -35.6203 | -35.3584
24.1057 | -35.6208 | -35.3644
24.1118 | -35.6212 | -35.3704
24.1179 | -35.6217 | -35.3763
24.124 | -35.6221 | -35.3822
24.1301 | -35.6225 | -35.3879
24.1362 | -35.6229 | -35.3937
24.1424 | -35.6233 | -35.3995
24.1485 | -35.6237 | -35.4054
24.1546 | -35.6241 | -35.4113
24.1607 | -35.6246 | -35.4172
24.1668 | -35.625 | -35.4229
24.1729 | -35.6254 | -35.429
24.179 | -35.6258 | -35.4348
24.1851 | -35.6263 | -35.4407
24.1912 | -35.6267 | -35.4466
24.1973 | -35.6271 | -35.4526
24.2034 | -35.6275 | -35.4586
24.2095 | -35.628 | -35.4649
24.2156 | -35.6284 | -35.4709
24.2217 | -35.6289 | -35.477
24.2278 | -35.6293 | -35.4831
24.2339 | -35.6297 | -35.4893
24.24 | -35.6301 | -35.4949
24.2461 | -35.6306 | -35.5007
24.2522 | -35.631 | -35.5068
24.2583 | -35.6314 | -35.5126
24.2644 | -35.6319 | -35.5187
24.2705 | -35.6323 | -35.5244
24.2766 | -35.6327 | -35.5302
24.2827 | -35.6332 | -35.536
24.2889 | -35.6336 | -35.5418
24.295 | -35.634 | -35.5477
24.3011 | -35.6344 | -35.5535
24.3072 | -35.6348 | -35.5591
24.3133 | -35.6353 | -35.5648
24.3194 | -35.6357 | -35.5708
24.3255 | -35.6361 | -35.5766
24.3316 | -35.6366 | -35.5825
24.3377 | -35.637 | -35.5883
24.3438 | -35.6374 | -35.5941
24.3499 | -35.6378 | -35.5998
24.356 | -35.6382 | -35.6058
24.3621 | -35.6386 | -35.6115
24.3682 | -35.6391 | -35.6174
24.3743 | -35.6395 | -35.6234
24.3804 | -35.6399 | -35.629
24.3865 | -35.6404 | -35.6347
24.3926 | -35.6408 | -35.6406
24.3987 | -35.6412 | -35.6465
24.4048 | -35.6417 | -35.6524
24.4109 | -35.6421 | -35.658
24.417 | -35.6425 | -35.6637
24.4231 | -35.6429 | -35.6696
24.4293 | -35.6434 | -35.6754
24.4354 | -35.6438 | -35.6813
24.4415 | -35.6442 | -35.6871
24.4476 | -35.6447 | -35.6929
24.4537 | -35.6451 | -35.6986
24.4598 | -35.6455 | -35.7044
24.4659 | -35.646 | -35.7104
24.472 | -35.6464 | -35.7162
24.4781 | -35.6468 | -35.722
24.4842 | -35.6472 | -35.7278
24.4903 | -35.6476 | -35.7337
24.4964 | -35.648 | -35.7395
24.5025 | -35.6484 | -35.7455
24.5086 | -35.6489 | -35.7513
24.5147 | -35.6493 | -35.7571
24.5208 | -35.6497 | -35.7627
24.5269 | -35.6502 | -35.7685
24.533 | -35.6506 | -35.7745
24.5391 | -35.651 | -35.7803
24.5452 | -35.6515 | -35.7862
24.5513 | -35.6519 | -35.7921
24.5574 | -35.6523 | -35.7978
24.5635 | -35.6527 | -35.8035
24.5696 | -35.6532 | -35.8094
24.5758 | -35.6536 | -35.8152
24.5819 | -35.654 | -35.821
24.588 | -35.6544 | -35.8266
24.5941 | -35.6549 | -35.8325
24.6002 | -35.6553 | -35.8383
24.6063 | -35.6557 | -35.8442
24.6124 | -35.6561 | -35.85
24.6185 | -35.6566 | -35.856
24.6246 | -35.6569 | -35.8617
24.6307 | -35.6574 | -35.8675
24.6368 | -35.6578 | -35.8735
24.6429 | -35.6582 | -35.8792
24.649 | -35.6586 | -35.885
24.6551 | -35.6591 | -35.8909
24.6612 | -35.6595 | -35.8966
24.6673 | -35.6599 | -35.9025
24.6734 | -35.6603 | -35.9083
24.6795 | -35.6607 | -35.9143
24.6856 | -35.6611 | -35.9202
24.6917 | -35.6616 | -35.9264
24.6978 | -35.662 | -35.9319
24.7039 | -35.6624 | -35.9379
24.71 | -35.6628 | -35.9439
24.7162 | -35.6632 | -35.9497
24.7223 | -35.6637 | -35.9555
24.7284 | -35.6641 | -35.9613
24.7345 | -35.6645 | -35.9669
24.7406 | -35.6649 | -35.9727
24.7467 | -35.6653 | -35.9787
24.7528 | -35.6657 | -35.9845
24.7589 | -35.6661 | -35.9901
24.765 | -35.6665 | -35.9957
24.7711 | -35.6669 | -36.0015
24.7772 | -35.6674 | -36.0071
24.7833 | -35.6678 | -36.0128
24.7894 | -35.6682 | -36.0188
24.7955 | -35.6686 | -36.0247
24.8016 | -35.669 | -36.0302
24.8077 | -35.6695 | -36.0359
24.8138 | -35.6699 | -36.0416
24.8199 | -35.6703 | -36.0475
24.826 | -35.6707 | -36.0533
24.8321 | -35.6711 | -36.0592
24.8382 | -35.6715 | -36.0648
24.8443 | -35.6719 | -36.0703
24.8504 | -35.6723 | -36.0763
24.8565 | -35.6727 | -36.0822
24.8627 | -35.6731 | -36.0879
24.8688 | -35.6735 | -36.0938
24.8749 | -35.6739 | -36.0996
24.881 | -35.6744 | -36.1055
24.8871 | -35.6748 | -36.1116
24.8932 | -35.6752 | -36.1175
24.8993 | -35.6756 | -36.1232
24.9054 | -35.676 | -36.129
24.9115 | -35.6764 | -36.1349
24.9176 | -35.6769 | -36.1408
24.9237 | -35.6773 | -36.1465
24.9298 | -35.6777 | -36.1523
24.9359 | -35.6781 | -36.158
24.942 | -35.6785 | -36.1636
24.9481 | -35.6789 | -36.1694
24.9542 | -35.6793 | -36.1753
24.9603 | -35.6797 | -36.1811
24.9664 | -35.6801 | -36.1869
24.9725 | -35.6805 | -36.1928
24.9786 | -35.6809 | -36.1984
24.9847 | -35.6813 | -36.2042
24.9908 | -35.6817 | -36.21
24.9969 | -35.6821 | -36.2158
25.0031 | -35.6825 | -36.2216
25.0092 | -35.6829 | -36.2273
25.0153 | -35.6833 | -36.2329
25.0214 | -35.6837 | -36.2388
25.0275 | -35.6841 | -36.2445
25.0336 | -35.6845 | -36.2503
25.0397 | -35.6849 | -36.2561
25.0458 | -35.6853 | -36.2616
25.0519 | -35.6857 | -36.2674
25.058 | -35.6861 | -36.2732
25.0641 | -35.6865 | -36.2792
25.0702 | -35.6869 | -36.2849
25.0763 | -35.6874 | -36.2907
25.0824 | -35.6878 | -36.2964
25.0885 | -35.6882 | -36.3022
25.0946 | -35.6886 | -36.308
25.1007 | -35.689 | -36.3139
25.1068 | -35.6894 | -36.3197
25.1129 | -35.6898 | -36.3254
25.119 | -35.6902 | -36.3309
25.1251 | -35.6906 | -36.3368
25.1312 | -35.6911 | -36.3426
25.1373 | -35.6915 | -36.3486
25.1435 | -35.6919 | -36.3545
25.1496 | -35.6923 | -36.36
25.1557 | -35.6926 | -36.3658
25.1618 | -35.6931 | -36.3714
25.1679 | -35.6935 | -36.3773
25.174 | -35.6939 | -36.3832
25.1801 | -35.6943 | -36.3889
25.1862 | -35.6946 | -36.3947
25.1923 | -35.695 | -36.4004
25.1984 | -35.6954 | -36.4062
25.2045 | -35.6958 | -36.4119
25.2106 | -35.6962 | -36.4178
25.2167 | -35.6966 | -36.4236
25.2228 | -35.697 | -36.4295
25.2289 | -35.6974 | -36.4352
25.235 | -35.6978 | -36.441
25.2411 | -35.6983 | -36.4469
25.2472 | -35.6986 | -36.4527
25.2533 | -35.699 | -36.4586
25.2594 | -35.6994 | -36.4642
25.2655 | -35.6998 | -36.47
25.2716 | -35.7002 | -36.4758
25.2777 | -35.7006 | -36.4814
25.2838 | -35.701 | -36.4871
25.29 | -35.7014 | -36.4927
25.2961 | -35.7018 | -36.4983
25.3022 | -35.7022 | -36.504
25.3083 | -35.7026 | -36.5098
25.3144 | -35.703 | -36.5155
25.3205 | -35.7034 | -36.5212
25.3266 | -35.7038 | -36.5269
25.3327 | -35.7042 | -36.5327
25.3388 | -35.7046 | -36.5385
25.3449 | -35.705 | -36.5442
25.351 | -35.7054 | -36.55
25.3571 | -35.7057 | -36.5559
25.3632 | -35.7061 | -36.5617
25.3693 | -35.7065 | -36.5673
25.3754 | -35.7069 | -36.5731
25.3815 | -35.7073 | -36.579
25.3876 | -35.7077 | -36.5849
25.3937 | -35.7081 | -36.5907
25.3998 | -35.7085 | -36.5961
25.4059 | -35.7089 | -36.6019
25.412 | -35.7093 | -36.6078
25.4181 | -35.7097 | -36.6136
25.4242 | -35.7101 | -36.6193
25.4304 | -35.7105 | -36.6251
25.4365 | -35.7109 | -36.6306
25.4426 | -35.7113 | -36.6364
25.4487 | -35.7116 | -36.6421
25.4548 | -35.712 | -36.6479
25.4609 | -35.7124 | -36.6535
25.467 | -35.7128 | -36.6591
25.4731 | -35.7132 | -36.6648
25.4792 | -35.7136 | -36.6706
25.4853 | -35.714 | -36.6765
25.4914 | -35.7145 | -36.6822
25.4975 | -35.7149 | -36.6879
25.5036 | -35.7153 | -36.6936
25.5097 | -35.7156 | -36.6992
25.5158 | -35.716 | -36.7051
25.5219 | -35.7164 | -36.711
25.528 | -35.7168 | -36.7168
25.5341 | -35.7172 | -36.7223
25.5402 | -35.7176 | -36.728
25.5463 | -35.718 | -36.7338
25.5524 | -35.7184 | -36.7393
25.5585 | -35.7188 | -36.7449
25.5646 | -35.7192 | -36.7508
25.5707 | -35.7196 | -36.7566
25.5769 | -35.72 | -36.762
25.583 | -35.7204 | -36.7676
25.5891 | -35.7208 | -36.7734
25.5952 | -35.7212 | -36.7793
25.6013 | -35.7216 | -36.7849
25.6074 | -35.722 | -36.7906
25.6135 | -35.7224 | -36.7963
25.6196 | -35.7227 | -36.8021
25.6257 | -35.7231 | -36.8078
25.6318 | -35.7235 | -36.8135
25.6379 | -35.7239 | -36.8194
25.644 | -35.7242 | -36.8251
25.6501 | -35.7246 | -36.8308
25.6562 | -35.725 | -36.8365
25.6623 | -35.7254 | -36.8421
25.6684 | -35.7258 | -36.8478
25.6745 | -35.7262 | -36.8536
25.6806 | -35.7266 | -36.8593
25.6867 | -35.727 | -36.865
25.6928 | -35.7273 | -36.8707
25.6989 | -35.7277 | -36.8765
25.705 | -35.7281 | -36.8822
25.7111 | -35.7285 | -36.8878
25.7173 | -35.7289 | -36.8936
25.7234 | -35.7293 | -36.8994
25.7295 | -35.7297 | -36.905
25.7356 | -35.73 | -36.9108
25.7417 | -35.7304 | -36.9165
25.7478 | -35.7308 | -36.9221
25.7539 | -35.7312 | -36.9278
25.76 | -35.7315 | -36.9337
25.7661 | -35.7319 | -36.9393
25.7722 | -35.7323 | -36.9451
25.7783 | -35.7327 | -36.9508
25.7844 | -35.733 | -36.9564
25.7905 | -35.7334 | -36.9621
25.7966 | -35.7338 | -36.9677
25.8027 | -35.7341 | -36.9734
25.8088 | -35.7345 | -36.9793
25.8149 | -35.7349 | -36.985
25.821 | -35.7353 | -36.9907
25.8271 | -35.7356 | -36.9964
25.8332 | -35.736 | -37.002
25.8393 | -35.7364 | -37.0077
25.8454 | -35.7368 | -37.0133
25.8515 | -35.7372 | -37.0188
25.8576 | -35.7375 | -37.0243
25.8638 | -35.7379 | -37.0303
25.8699 | -35.7383 | -37.036
25.876 | -35.7387 | -37.0418
25.8821 | -35.7391 | -37.0474
25.8882 | -35.7394 | -37.053
25.8943 | -35.7398 | -37.0587
25.9004 | -35.7402 | -37.0644
25.9065 | -35.7405 | -37.0703
25.9126 | -35.7409 | -37.0758
25.9187 | -35.7413 | -37.0815
25.9248 | -35.7417 | -37.087
25.9309 | -35.742 | -37.0926
25.937 | -35.7424 | -37.0982
25.9431 | -35.7428 | -37.104
25.9492 | -35.7432 | -37.1096
25.9553 | -35.7436 | -37.1154
25.9614 | -35.7439 | -37.121
25.9675 | -35.7443 | -37.1269
25.9736 | -35.7446 | -37.1327
25.9797 | -35.745 | -37.1386
25.9858 | -35.7454 | -37.1442
25.9919 | -35.7458 | -37.1501
25.998 | -35.7461 | -37.1558
26.0042 | -35.7465 | -37.1615
26.0103 | -35.7469 | -37.1673
26.0164 | -35.7473 | -37.1731
26.0225 | -35.7477 | -37.1787
26.0286 | -35.7481 | -37.1844
26.0347 | -35.7484 | -37.19
26.0408 | -35.7488 | -37.1959
26.0469 | -35.7491 | -37.2017
26.053 | -35.7495 | -37.2074
26.0591 | -35.7499 | -37.2131
26.0652 | -35.7503 | -37.2187
26.0713 | -35.7507 | -37.2245
26.0774 | -35.751 | -37.2302
26.0835 | -35.7514 | -37.2358
26.0896 | -35.7518 | -37.2416
26.0957 | -35.7522 | -37.2474
26.1018 | -35.7525 | -37.253
26.1079 | -35.7529 | -37.2587
26.114 | -35.7533 | -37.2643
26.1201 | -35.7536 | -37.27
26.1262 | -35.754 | -37.2755
26.1323 | -35.7544 | -37.2812
26.1384 | -35.7547 | -37.2869
26.1445 | -35.7551 | -37.2927
26.1507 | -35.7555 | -37.2985
26.1568 | -35.7558 | -37.3041
26.1629 | -35.7562 | -37.3098
26.169 | -35.7566 | -37.3154
26.1751 | -35.7569 | -37.3211
26.1812 | -35.7573 | -37.3268
26.1873 | -35.7577 | -37.3325
26.1934 | -35.7581 | -37.3381
26.1995 | -35.7584 | -37.3438
26.2056 | -35.7588 | -37.3495
26.2117 | -35.7591 | -37.3551
26.2178 | -35.7595 | -37.3607
26.2239 | -35.7599 | -37.3664
26.23 | -35.7602 | -37.3722
26.2361 | -35.7606 | -37.3779
26.2422 | -35.761 | -37.3835
26.2483 | -35.7613 | -37.3891
26.2544 | -35.7617 | -37.3947
26.2605 | -35.7621 | -37.4002
26.2666 | -35.7625 | -37.4058
26.2727 | -35.7628 | -37.4115
26.2788 | -35.7632 | -37.417
26.2849 | -35.7636 | -37.4227
26.2911 | -35.7639 | -37.4284
26.2972 | -35.7643 | -37.4341
26.3033 | -35.7647 | -37.4398
26.3094 | -35.765 | -37.4455
26.3155 | -35.7654 | -37.4511
26.3216 | -35.7657 | -37.457
26.3277 | -35.7661 | -37.4628
26.3338 | -35.7665 | -37.4683
26.3399 | -35.7668 | -37.4739
26.346 | -35.7672 | -37.4795
26.3521 | -35.7676 | -37.4853
26.3582 | -35.768 | -37.491
26.3643 | -35.7683 | -37.4965
26.3704 | -35.7687 | -37.5023
26.3765 | -35.7691 | -37.5079
26.3826 | -35.7694 | -37.5134
26.3887 | -35.7698 | -37.5191
26.3948 | -35.7701 | -37.5248
26.4009 | -35.7705 | -37.5306
26.407 | -35.7709 | -37.536
26.4131 | -35.7712 | -37.5417
26.4192 | -35.7716 | -37.5472
26.4253 | -35.772 | -37.5528
26.4314 | -35.7723 | -37.5585
26.4376 | -35.7727 | -37.5642
26.4437 | -35.7731 | -37.5698
26.4498 | -35.7734 | -37.5756
26.4559 | -35.7738 | -37.5812
26.462 | -35.7742 | -37.587
26.4681 | -35.7745 | -37.5926
26.4742 | -35.7749 | -37.5984
26.4803 | -35.7752 | -37.6041
26.4864 | -35.7756 | -37.6095
26.4925 | -35.776 | -37.6152
26.4986 | -35.7764 | -37.6208
26.5047 | -35.7768 | -37.6265
26.5108 | -35.7771 | -37.6321
26.5169 | -35.7775 | -37.6379
26.523 | -35.7778 | -37.6435
26.5291 | -35.7782 | -37.6491
26.5352 | -35.7786 | -37.6548
26.5413 | -35.7789 | -37.6605
26.5474 | -35.7793 | -37.6662
26.5535 | -35.7797 | -37.6719
26.5596 | -35.78 | -37.6774
26.5657 | -35.7804 | -37.6831
26.5718 | -35.7808 | -37.6887
26.578 | -35.7811 | -37.6943
26.5841 | -35.7815 | -37.6998
26.5902 | -35.7819 | -37.7055
26.5963 | -35.7822 | -37.7109
26.6024 | -35.7826 | -37.7166
26.6085 | -35.783 | -37.7224
26.6146 | -35.7833 | -37.7281
26.6207 | -35.7837 | -37.7339
26.6268 | -35.7841 | -37.7396
26.6329 | -35.7844 | -37.7452
26.639 | -35.7848 | -37.7509
26.6451 | -35.7851 | -37.7566
26.6512 | -35.7855 | -37.7621
26.6573 | -35.7858 | -37.7675
26.6634 | -35.7862 | -37.773
26.6695 | -35.7865 | -37.7788
26.6756 | -35.7869 | -37.7844
26.6817 | -35.7872 | -37.7902
26.6878 | -35.7876 | -37.7959
26.6939 | -35.788 | -37.8015
26.7 | -35.7883 | -37.8072
26.7061 | -35.7886 | -37.813
26.7122 | -35.789 | -37.8186
26.7183 | -35.7894 | -37.8241
26.7245 | -35.7897 | -37.8297
26.7306 | -35.7901 | -37.8353
26.7367 | -35.7904 | -37.8408
26.7428 | -35.7908 | -37.8464
26.7489 | -35.7911 | -37.8522
26.755 | -35.7915 | -37.8577
26.7611 | -35.7918 | -37.8633
26.7672 | -35.7922 | -37.8688
26.7733 | -35.7925 | -37.8743
26.7794 | -35.7929 | -37.88
26.7855 | -35.7933 | -37.8856
26.7916 | -35.7936 | -37.8912
26.7977 | -35.794 | -37.8968
26.8038 | -35.7943 | -37.9025
26.8099 | -35.7947 | -37.908
26.816 | -35.7951 | -37.9137
26.8221 | -35.7954 | -37.9193
26.8282 | -35.7957 | -37.9247
26.8343 | -35.7961 | -37.9303
26.8404 | -35.7964 | -37.9359
26.8465 | -35.7968 | -37.9416
26.8526 | -35.7971 | -37.9472
26.8587 | -35.7975 | -37.9527
26.8649 | -35.7978 | -37.9583
26.871 | -35.7982 | -37.964
26.8771 | -35.7986 | -37.9695
26.8832 | -35.7989 | -37.975
26.8893 | -35.7993 | -37.9807
26.8954 | -35.7996 | -37.9863
26.9015 | -35.8 | -37.992
26.9076 | -35.8003 | -37.9976
26.9137 | -35.8007 | -38.0032
26.9198 | -35.801 | -38.0091
26.9259 | -35.8014 | -38.0147
26.932 | -35.8017 | -38.0202
26.9381 | -35.8021 | -38.0257
26.9442 | -35.8024 | -38.0313
26.9503 | -35.8028 | -38.0369
26.9564 | -35.8031 | -38.0424
26.9625 | -35.8035 | -38.0482
26.9686 | -35.8038 | -38.0538
26.9747 | -35.8042 | -38.0596
26.9808 | -35.8045 | -38.0651
26.9869 | -35.8049 | -38.0706
26.993 | -35.8053 | -38.0762
26.9991 | -35.8056 | -38.0819
27.0052 | -35.8059 | -38.0875
27.0114 | -35.8063 | -38.093
27.0175 | -35.8066 | -38.0982
27.0236 | -35.807 | -38.1037
27.0297 | -35.8073 | -38.1093
27.0358 | -35.8077 | -38.1147
27.0419 | -35.8081 | -38.1205
27.048 | -35.8084 | -38.1263
27.0541 | -35.8088 | -38.1319
27.0602 | -35.8091 | -38.1375
27.0663 | -35.8095 | -38.1432
27.0724 | -35.8098 | -38.1488
27.0785 | -35.8102 | -38.1546
27.0846 | -35.8105 | -38.16
27.0907 | -35.8109 | -38.1656
27.0968 | -35.8112 | -38.1712
27.1029 | -35.8115 | -38.1768
27.109 | -35.8119 | -38.1823
27.1151 | -35.8123 | -38.1878
27.1212 | -35.8126 | -38.1932
27.1273 | -35.813 | -38.1988
27.1334 | -35.8133 | -38.2044
27.1395 | -35.8137 | -38.2099
27.1456 | -35.814 | -38.2154
27.1518 | -35.8144 | -38.221
27.1579 | -35.8147 | -38.2265
27.164 | -35.8151 | -38.2319
27.1701 | -35.8155 | -38.2374
27.1762 | -35.8158 | -38.2429
27.1823 | -35.8162 | -38.2484
27.1884 | -35.8165 | -38.254
27.1945 | -35.8169 | -38.2595
27.2006 | -35.8172 | -38.2649
27.2067 | -35.8176 | -38.2705
27.2128 | -35.8179 | -38.276
27.2189 | -35.8183 | -38.2817
27.225 | -35.8186 | -38.2871
27.2311 | -35.8189 | -38.2925
27.2372 | -35.8193 | -38.2982
27.2433 | -35.8196 | -38.3038
27.2494 | -35.82 | -38.3094
27.2555 | -35.8203 | -38.3151
27.2616 | -35.8206 | -38.3204
27.2677 | -35.821 | -38.3258
27.2738 | -35.8213 | -38.3314
27.2799 | -35.8217 | -38.337
27.286 | -35.822 | -38.3427
27.2921 | -35.8224 | -38.3483
27.2983 | -35.8227 | -38.3538
27.3044 | -35.8231 | -38.3594
27.3105 | -35.8234 | -38.3649
27.3166 | -35.8238 | -38.3703
27.3227 | -35.8241 | -38.3759
27.3288 | -35.8244 | -38.3814
27.3349 | -35.8248 | -38.3867
27.341 | -35.8251 | -38.3924
27.3471 | -35.8254 | -38.3978
27.3532 | -35.8258 | -38.4033
27.3593 | -35.8261 | -38.4088
27.3654 | -35.8265 | -38.4145
27.3715 | -35.8268 | -38.4199
27.3776 | -35.8272 | -38.4255
27.3837 | -35.8275 | -38.4311
27.3898 | -35.8278 | -38.4365
27.3959 | -35.8282 | -38.4421
27.402 | -35.8285 | -38.4473
27.4081 | -35.8288 | -38.4529
27.4142 | -35.8292 | -38.4585
27.4203 | -35.8295 | -38.4642
27.4264 | -35.8298 | -38.4697
27.4325 | -35.8302 | -38.4753
27.4387 | -35.8305 | -38.4807
27.4448 | -35.8308 | -38.4862
27.4509 | -35.8312 | -38.4918
27.457 | -35.8315 | -38.4974
27.4631 | -35.8318 | -38.503
27.4692 | -35.8322 | -38.5085
27.4753 | -35.8325 | -38.5139
27.4814 | -35.8329 | -38.5195
27.4875 | -35.8332 | -38.5251
27.4936 | -35.8335 | -38.5308
27.4997 | -35.8339 | -38.5364
27.5058 | -35.8342 | -38.5419
27.5119 | -35.8345 | -38.5474
27.518 | -35.8349 | -38.553
27.5241 | -35.8352 | -38.5587
27.5302 | -35.8356 | -38.5643
27.5363 | -35.8359 | -38.5699
27.5424 | -35.8362 | -38.5752
27.5485 | -35.8366 | -38.5807
27.5546 | -35.8369 | -38.5862
27.5607 | -35.8373 | -38.5917
27.5668 | -35.8376 | -38.5972
27.5729 | -35.8379 | -38.6025
27.5791 | -35.8383 | -38.608
27.5852 | -35.8386 | -38.6134
27.5913 | -35.839 | -38.6189
27.5974 | -35.8393 | -38.6246
27.6035 | -35.8396 | -38.6302
27.6096 | -35.84 | -38.6356
27.6157 | -35.8403 | -38.641
27.6218 | -35.8407 | -38.6467
27.6279 | -35.841 | -38.6523
27.634 | -35.8413 | -38.6578
27.6401 | -35.8416 | -38.6632
27.6462 | -35.842 | -38.6689
27.6523 | -35.8423 | -38.6743
27.6584 | -35.8427 | -38.6799
27.6645 | -35.843 | -38.6855
27.6706 | -35.8433 | -38.6911
27.6767 | -35.8437 | -38.6967
27.6828 | -35.844 | -38.702
27.6889 | -35.8443 | -38.7076
27.695 | -35.8447 | -38.713
27.7011 | -35.845 | -38.7184
27.7072 | -35.8453 | -38.724
27.7133 | -35.8456 | -38.7295
27.7194 | -35.846 | -38.7349
27.7256 | -35.8463 | -38.7404
27.7317 | -35.8467 | -38.7458
27.7378 | -35.847 | -38.7513
27.7439 | -35.8474 | -38.7567
27.75 | -35.8477 | -38.7624
27.7561 | -35.848 | -38.7675
27.7622 | -35.8484 | -38.773
27.7683 | -35.8487 | -38.7784
27.7744 | -35.8491 | -38.784
27.7805 | -35.8494 | -38.7894
27.7866 | -35.8498 | -38.7951
27.7927 | -35.8501 | -38.8003
27.7988 | -35.8504 | -38.8059
27.8049 | -35.8508 | -38.8113
27.811 | -35.8511 | -38.8168
27.8171 | -35.8515 | -38.8222
27.8232 | -35.8518 | -38.8276
27.8293 | -35.8522 | -38.833
27.8354 | -35.8525 | -38.8386
27.8415 | -35.8528 | -38.844
27.8476 | -35.8532 | -38.8497
27.8537 | -35.8535 | -38.8553
27.8598 | -35.8538 | -38.8605
27.866 | -35.8541 | -38.866
27.8721 | -35.8545 | -38.8715
27.8782 | -35.8548 | -38.877
27.8843 | -35.8551 | -38.8824
27.8904 | -35.8555 | -38.888
27.8965 | -35.8558 | -38.8933
27.9026 | -35.8561 | -38.8988
27.9087 | -35.8565 | -38.9044
27.9148 | -35.8568 | -38.9099
27.9209 | -35.8571 | -38.9154
27.927 | -35.8575 | -38.9208
27.9331 | -35.8578 | -38.9261
27.9392 | -35.8581 | -38.9315
27.9453 | -35.8585 | -38.9369
27.9514 | -35.8588 | -38.9424
27.9575 | -35.8591 | -38.9479
27.9636 | -35.8594 | -38.9535
27.9697 | -35.8597 | -38.9588
27.9758 | -35.8601 | -38.9643
27.9819 | -35.8604 | -38.9697
27.988 | -35.8607 | -38.9751
27.9941 | -35.861 | -38.9805
28.0002 | -35.8614 | -38.9857
28.0063 | -35.8617 | -38.991
28.0125 | -35.862 | -38.9965
28.0186 | -35.8624 | -39.002
28.0247 | -35.8627 | -39.0074
28.0308 | -35.8631 | -39.0128
28.0369 | -35.8634 | -39.018
28.043 | -35.8637 | -39.0235
28.0491 | -35.8641 | -39.0289
28.0552 | -35.8644 | -39.0343
28.0613 | -35.8647 | -39.0396
28.0674 | -35.8651 | -39.0449
28.0735 | -35.8654 | -39.0501

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.606 | -44.0105
21.4908 | -35.6545 | -35.846
16.4948 | -35.0361 | -28.3369
12.6602 | -34.6684 | -21.9878
9.7171 | -34.4495 | -17.0314
7.4582 | -34.2987 | -13.0165
5.7244 | -34.1843 | -9.7758
4.3936 | -34.1094 | -7.1606
3.3722 | -34.084 | -5.0979
2.5883 | -34.0492 | -3.5378
1.9866 | -34.0441 | -2.2905
1.5248 | -34.0634 | -1.2786
1.1703 | -34.0702 | -0.62466
1 | -34.0692 | -0.31538
0.89822 | -34.0509 | -0.041733
0.68941 | -34.0246 | 0.53623
0.52919 | -34.0196 | 0.97524
0.40612 | -34.0046 | 1.5862
0.31171 | -34.0095 | 2.0228
0.23928 | -34.0077 | 2.7414
0.18358 | -34.0205 | 3.1615
0.14091 | -34.0399 | 4.0273
0.10819 | -34.0482 | 5.184
0.083017 | -34.1076 | 6.6966
0.063705 | -34.1755 | 8.5974
0.048923 | -34.2166 | 11.0667
0.037479 | -34.336 | 14.2969
0.028801 | -34.5697 | 18.3961
0.022078 | -34.8885 | 23.7857
0.017023 | -35.3751 | 30.2909
0.013018 | -36.1721 | 38.0126
0.010014 | -37.2888 | 45.9292

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003735 | -6.2312
0.00747 | -27.3996
0.011205 | -36.522
0.01494 | -41.1057
0.018675 | -44.852
0.02241 | -46.2476
0.026145 | -47.5445
0.02988 | -48.2797
0.033615 | -50.9038
0.03735 | -52.5469
0.041085 | -53.5848
0.04482 | -54.4866
0.048555 | -54.1814
0.05229 | -54.8674
0.056025 | -56.4794
0.05976 | -57.7837
0.063495 | -59.0583
0.06723 | -58.504
0.070965 | -59.41
0.0747 | -59.653
0.078435 | -60.4686
0.08217 | -61.5094
0.085905 | -61.3468
0.08964 | -62.0303
0.093375 | -62.7591
0.09711 | -63.6973
0.10085 | -63.6238
0.10458 | -63.4781
0.10831 | -65.07
0.11205 | -65.686
0.11578 | -65.2734
0.11952 | -66.3161
0.12326 | -66.6935
0.12699 | -66.9967
0.13073 | -66.7342
0.13446 | -66.707
0.1382 | -68.2665
0.14193 | -68.9201
0.14566 | -68.8223
0.1494 | -69.0242
0.15313 | -68.7301
0.15687 | -68.3244
0.1606 | -68.3217
0.16434 | -68.8231
0.16808 | -69.5523
0.17181 | -69.7602
0.17555 | -70.6542
0.17928 | -70.8423
0.18302 | -71.2152
0.18675 | -71.1118
0.19048 | -71.0875
0.19422 | -70.9272
0.19795 | -70.98
0.20169 | -71.6117
0.20542 | -70.5218
0.20916 | -71.3901
0.2129 | -71.4514
0.21663 | -71.028
0.22037 | -71.3856
0.2241 | -72.4352
0.22784 | -73.2356
0.23157 | -73.0038
0.2353 | -72.9673
0.23904 | -72.7934
0.24277 | -72.0279
0.24651 | -72.9446
0.25024 | -73.8155
0.25398 | -73.7034
0.25772 | -73.8838
0.26145 | -73.3061
0.26519 | -73.4582
0.26892 | -74.8011
0.27265 | -75.3229
0.27639 | -74.6236
0.28013 | -74.0243
0.28386 | -74.827
0.28759 | -75.4752
0.29133 | -75.1838
0.29507 | -75.4197
0.2988 | -76.5908
0.30253 | -75.6934
0.30627 | -75.4105
0.31 | -74.8816
0.31374 | -75.5619
0.31748 | -76.1552
0.32121 | -75.7965
0.32494 | -75.9279
0.32868 | -76.7421
0.33242 | -76.4725
0.33615 | -75.6958
0.33988 | -76.2466
0.34362 | -75.7579
0.34736 | -76.904
0.35109 | -77.8786
0.35483 | -77.4411
0.35856 | -77.1391
0.36229 | -77.3825
0.36603 | -77.3296
0.36977 | -76.2917
0.3735 | -76.3569
0.37723 | -77.8214
0.38097 | -79.5635
0.38471 | -78.5758
0.38844 | -77.644
0.39217 | -77.7637
0.39591 | -77.9012
0.39964 | -79.2274
0.40338 | -78.012
0.40712 | -78.2116
0.41085 | -78.3284
0.41458 | -78.6431
0.41832 | -79.1931
0.42206 | -79.1172
0.42579 | -79.3429
0.42952 | -79.1208
0.43326 | -78.8563
0.437 | -79.0641
0.44073 | -79.1496
0.44446 | -79.5117
0.4482 | -80.1723
0.45193 | -80.2593
0.45567 | -80.245
0.45941 | -79.5232
0.46314 | -79.9122
0.46687 | -79.8356
0.47061 | -80.3423
0.47435 | -80.8052
0.47808 | -80.046
0.48181 | -79.6745
0.48555 | -80.4053
0.48929 | -80.3599
0.49302 | -80.1033
0.49676 | -79.9037
0.50049 | -79.5887
0.50423 | -79.4662
0.50796 | -79.5899
0.5117 | -80.3298
0.51543 | -82.3064
0.51916 | -82.4432
0.5229 | -81.426
0.52663 | -80.5997
0.53037 | -80.6175
0.53411 | -81.2315
0.53784 | -81.8889
0.54158 | -81.868
0.54531 | -81.5632
0.54905 | -81.8153
0.55278 | -81.958
0.55651 | -81.7885
0.56025 | -81.613
0.56398 | -81.3567
0.56772 | -81.2589
0.57146 | -81.4328
0.57519 | -81.5667
0.57893 | -81.8864
0.58266 | -82.509
0.58639 | -82.447
0.59013 | -82.2142
0.59386 | -82.6045
0.5976 | -82.956
0.60133 | -83.2029
0.60507 | -83.3312
0.60881 | -83.0642
0.61254 | -82.4533
0.61628 | -82.147
0.62001 | -82.5923
0.62374 | -82.7445
0.62748 | -82.1981
0.63121 | -82.0272
0.63495 | -82.2115
0.63868 | -82.431
0.64242 | -82.9253
0.64616 | -83.4773
0.64989 | -84.1126
0.65363 | -84.4267
0.65736 | -84.0464
0.66109 | -83.7421
0.66483 | -83.619
0.66856 | -83.2458
0.6723 | -82.9023
0.67604 | -82.8498
0.67977 | -83.1819
0.68351 | -83.1972
0.68724 | -82.7393
0.69098 | -82.4715
0.69471 | -82.701
0.69844 | -82.8783
0.70218 | -83.1028
0.70591 | -83.4361
0.70965 | -83.9253
0.71339 | -84.4276
0.71712 | -84.6879
0.72086 | -84.9458
0.72459 | -84.8555
0.72833 | -84.2947
0.73206 | -83.7646
0.73579 | -83.4715
0.73953 | -83.5129
0.74326 | -83.8387
0.747 | -83.9382
0.75074 | -83.6617
0.75447 | -83.9099
0.75821 | -84.5586
0.76194 | -84.6722
0.76567 | -84.5747
0.76941 | -84.8327
0.77314 | -85.0182
0.77688 | -85.1317
0.78061 | -85.1971
0.78435 | -85.3812
0.78809 | -85.1425
0.79182 | -84.8897
0.79556 | -85.2623
0.79929 | -85.3892
0.80302 | -84.8778
0.80676 | -84.3779
0.81049 | -84.3341
0.81423 | -84.6724
0.81797 | -85.1731
0.8217 | -85.8126
0.82544 | -86.2205
0.82917 | -85.9631
0.83291 | -85.6861
0.83664 | -85.8036
0.84037 | -86.0483
0.84411 | -86.0171
0.84784 | -86.0734
0.85158 | -86.4695
0.85532 | -86.5269
0.85905 | -85.9631
0.86279 | -85.0959
0.86652 | -84.428
0.87026 | -84.6756
0.87399 | -85.8225
0.87772 | -86.909
0.88146 | -87.0204
0.88519 | -86.7041
0.88893 | -86.4234
0.89267 | -86.2251
0.8964 | -86.1939
0.90014 | -86.0103
0.90387 | -85.8923
0.9076 | -86.0132
0.91134 | -86.3287
0.91507 | -86.4382
0.91881 | -86.3524
0.92254 | -86.2193
0.92628 | -86.0443
0.93002 | -86.2058
0.93375 | -86.855
0.93749 | -87.1362
0.94122 | -86.6238
0.94495 | -86.7883
0.94869 | -87.2659
0.95242 | -86.8163
0.95616 | -86.3381
0.9599 | -86.2701
0.96363 | -86.7523
0.96737 | -87.6635
0.9711 | -87.7922
0.97484 | -87.1743
0.97857 | -87.0961
0.9823 | -87.4382
0.98604 | -87.3897
0.98977 | -87.3105
0.99351 | -87.5346
0.99725 | -87.5864
1.001 | -87.3609
1.0047 | -87.1698
1.0085 | -86.9011
1.0122 | -86.8388
1.0159 | -87.0189
1.0197 | -87.2097
1.0234 | -87.7024
1.0271 | -88.4187
1.0309 | -88.6993
1.0346 | -87.9468
1.0383 | -87.4049
1.0421 | -87.776
1.0458 | -88.104
1.0495 | -88.1152
1.0533 | -88.1337
1.057 | -87.8797
1.0607 | -87.7966
1.0645 | -87.9924
1.0682 | -87.95
1.0719 | -88.0789
1.0757 | -87.9841
1.0794 | -87.8296
1.0832 | -87.753
1.0869 | -87.8465
1.0906 | -88.0641
1.0944 | -88.2195
1.0981 | -88.1059
1.1018 | -87.8945
1.1056 | -87.9262
1.1093 | -88.1053
1.113 | -88.0633
1.1168 | -87.8505
1.1205 | -87.7819
1.1242 | -87.9321
1.128 | -88.1119
1.1317 | -88.2675
1.1354 | -88.6298
1.1392 | -89.0158
1.1429 | -88.7991
1.1466 | -88.1891
1.1504 | -87.8657
1.1541 | -87.9571
1.1579 | -88.1768
1.1616 | -88.3221
1.1653 | -88.4411
1.1691 | -88.6434
1.1728 | -88.9469
1.1765 | -89.2423
1.1803 | -89.4137
1.184 | -89.4721
1.1877 | -89.4262
1.1915 | -89.2841
1.1952 | -89.132
1.1989 | -89.0139
1.2027 | -88.8366
1.2064 | -88.6036
1.2101 | -88.4666
1.2139 | -88.5301
1.2176 | -88.7803
1.2213 | -89.1199
1.2251 | -89.4545
1.2288 | -89.6447
1.2326 | -89.4309
1.2363 | -88.8181
1.24 | -88.3028
1.2438 | -88.2985
1.2475 | -88.7027
1.2512 | -88.9915
1.255 | -88.9082
1.2587 | -88.7946
1.2624 | -88.9281
1.2662 | -89.0995
1.2699 | -88.9573
1.2736 | -88.6589
1.2774 | -88.5232
1.2811 | -88.5805
1.2848 | -88.7804
1.2886 | -89.1088
1.2923 | -89.4658
1.296 | -89.633
1.2998 | -89.6127
1.3035 | -89.6373
1.3073 | -89.6678
1.311 | -89.5666
1.3147 | -89.3954
1.3185 | -89.2793
1.3222 | -89.2497
1.3259 | -89.3054
1.3297 | -89.4621
1.3334 | -89.6027
1.3371 | -89.6439
1.3409 | -89.7353
1.3446 | -89.9462
1.3483 | -90.1702
1.3521 | -90.1793
1.3558 | -89.7395
1.3595 | -89.014
1.3633 | -88.5092
1.367 | -88.563
1.3707 | -89.1007
1.3745 | -89.7006
1.3782 | -89.9557
1.382 | -89.8819
1.3857 | -89.7931
1.3894 | -89.9174
1.3932 | -90.1973
1.3969 | -90.3601
1.4006 | -90.3045
1.4044 | -90.1398
1.4081 | -89.9595
1.4118 | -89.8435
1.4156 | -89.8688
1.4193 | -90.1252
1.423 | -90.4896
1.4268 | -90.6576
1.4305 | -90.59
1.4342 | -90.4074
1.438 | -90.2222
1.4417 | -90.1102
1.4454 | -90.0457
1.4492 | -89.9879
1.4529 | -89.9605
1.4567 | -90.0647
1.4604 | -90.3441
1.4641 | -90.7004
1.4679 | -91.0675
1.4716 | -91.4236
1.4753 | -91.5905
1.4791 | -91.3867
1.4828 | -90.917
1.4865 | -90.3879
1.4903 | -90.0449
1.494 | -90.1195
1.4977 | -90.584
1.5015 | -91.096
1.5052 | -91.4451
1.5089 | -91.7613
1.5127 | -91.9263
1.5164 | -91.5624
1.5201 | -90.9513
1.5239 | -90.6543
1.5276 | -90.7007
1.5314 | -90.797
1.5351 | -90.8157
1.5388 | -90.897
1.5426 | -91.0742
1.5463 | -91.1768
1.55 | -91.1646
1.5538 | -91.163
1.5575 | -91.13
1.5612 | -90.8963
1.565 | -90.6207
1.5687 | -90.6096
1.5724 | -90.8756
1.5762 | -91.2102
1.5799 | -91.3736
1.5836 | -91.2577
1.5874 | -91.0788
1.5911 | -91.1253
1.5948 | -91.2765
1.5986 | -91.2662
1.6023 | -91.2236
1.6061 | -91.3862
1.6098 | -91.7003
1.6135 | -91.6259
1.6173 | -91.5385
1.621 | -91.4647
1.6247 | -91.5838
1.6285 | -91.9152
1.6322 | -92.3449
1.6359 | -92.6707
1.6397 | -92.7548
1.6434 | -92.6355
1.6471 | -92.442
1.6509 | -92.2642
1.6546 | -92.1255
1.6583 | -92.0285
1.6621 | -92.0079
1.6658 | -92.0842
1.6695 | -92.2254
1.6733 | -92.337
1.677 | -92.3438
1.6808 | -92.2544
1.6845 | -92.1338
1.6882 | -92.0322
1.692 | -91.9737
1.6957 | -91.991
1.6994 | -92.0972
1.7032 | -92.2463
1.7069 | -92.3613
1.7106 | -92.4239
1.7144 | -92.4534
1.7181 | -92.4659
1.7218 | -92.4642
1.7256 | -92.4718
1.7293 | -92.5486
1.733 | -92.7335
1.7368 | -92.9762
1.7405 | -93.0996
1.7442 | -92.9667
1.748 | -92.6637
1.7517 | -92.3748
1.7555 | -92.2547
1.7592 | -92.3536
1.7629 | -92.6053
1.7667 | -92.854
1.7704 | -92.9379
1.7741 | -92.8727
1.7779 | -92.7485
1.7816 | -92.5942
1.7853 | -92.4003
1.7891 | -92.2234
1.7928 | -92.176
1.7965 | -92.2913
1.8003 | -92.4893
1.804 | -92.6367
1.8077 | -92.6504
1.8115 | -92.577
1.8152 | -92.5221
1.8189 | -92.5741
1.8227 | -92.7397
1.8264 | -92.938
1.8302 | -93.0416
1.8339 | -92.9692
1.8376 | -92.7598
1.8414 | -92.5
1.8451 | -92.2926
1.8488 | -92.2184
1.8526 | -92.3087
1.8563 | -92.5174
1.86 | -92.7648
1.8638 | -92.9909
1.8675 | -93.1779
1.8712 | -93.3417
1.875 | -93.4937
1.8787 | -93.5574
1.8824 | -93.44
1.8862 | -93.1561
1.8899 | -92.8285
1.8936 | -92.5653
1.8974 | -92.3761
1.9011 | -92.2469
1.9049 | -92.2285
1.9086 | -92.3401
1.9123 | -92.4846
1.9161 | -92.5084
1.9198 | -92.3971
1.9235 | -92.3419
1.9273 | -92.4793
1.931 | -92.8023
1.9347 | -93.1894
1.9385 | -93.4861
1.9422 | -93.5987
1.9459 | -93.5157
1.9497 | -93.3305
1.9534 | -93.137
1.9571 | -92.9345
1.9609 | -92.7708
1.9646 | -92.7171
1.9683 | -92.8312
1.9721 | -93.0744
1.9758 | -93.306
1.9796 | -93.3838
1.9833 | -93.2891
1.987 | -93.1367
1.9908 | -93.0676
1.9945 | -93.0918
1.9982 | -93.105
2.002 | -92.988
2.0057 | -92.792
2.0094 | -92.6448
2.0132 | -92.6257
2.0169 | -92.7485
2.0206 | -93
2.0244 | -93.3096
2.0281 | -93.5932
2.0318 | -93.7858
2.0356 | -93.9
2.0393 | -93.9333
2.043 | -93.8485
2.0468 | -93.6784
2.0505 | -93.5321
2.0543 | -93.4452
2.058 | -93.3692
2.0617 | -93.221
2.0655 | -93.0365
2.0692 | -92.906
2.0729 | -92.8605
2.0767 | -92.8596
2.0804 | -92.8324
2.0841 | -92.8185
2.0879 | -92.9434
2.0916 | -93.2329
2.0953 | -93.5735
2.0991 | -93.7835
2.1028 | -93.7829
2.1065 | -93.5897
2.1103 | -93.3245
2.114 | -93.1395
2.1177 | -93.0744
2.1215 | -93.0837
2.1252 | -93.1716
2.129 | -93.3632
2.1327 | -93.5785
2.1364 | -93.6429
2.1402 | -93.4695
2.1439 | -93.1628
2.1476 | -92.9456
2.1514 | -92.9541
2.1551 | -93.2222
2.1588 | -93.4058
2.1626 | -93.5482
2.1663 | -93.5933
2.17 | -93.5622
2.1738 | -93.5244
2.1775 | -93.558
2.1812 | -93.6906
2.185 | -93.8884
2.1887 | -94.0892
2.1924 | -94.2326
2.1962 | -94.2988
2.1999 | -94.3173
2.2037 | -94.3411
2.2074 | -94.3732
2.2111 | -94.4001
2.2149 | -94.3914
2.2186 | -94.3358
2.2223 | -94.2281
2.2261 | -94.0758
2.2298 | -93.8947
2.2335 | -93.7595
2.2373 | -93.7057
2.241 | -93.7359
2.2447 | -93.8141
2.2485 | -93.897
2.2522 | -93.9387
2.2559 | -93.9288
2.2597 | -93.8686
2.2634 | -93.813
2.2671 | -93.8108
2.2709 | -93.8898
2.2746 | -94.0256
2.2784 | -94.1949
2.2821 | -94.3797
2.2858 | -94.5673
2.2896 | -94.7448
2.2933 | -94.8594
2.297 | -94.8527
2.3008 | -94.7196
2.3045 | -94.5107
2.3082 | -94.303
2.312 | -94.1387
2.3157 | -94.0498
2.3194 | -94.0463
2.3232 | -94.1114
2.3269 | -94.2265
2.3306 | -94.3746
2.3344 | -94.5628
2.3381 | -94.7852
2.3418 | -95.0065
2.3456 | -95.135
2.3493 | -95.1358
2.3531 | -95.0375
2.3568 | -94.9277
2.3605 | -94.9015
2.3643 | -94.9674
2.368 | -95.0829
2.3717 | -95.1491
2.3755 | -95.1272
2.3792 | -95.0457
2.3829 | -94.9481
2.3867 | -94.8536
2.3904 | -94.7553
2.3941 | -94.6786
2.3979 | -94.627
2.4016 | -94.601
2.4053 | -94.5827
2.4091 | -94.5542
2.4128 | -94.5212
2.4165 | -94.4515
2.4203 | -94.3565
2.424 | -94.2551
2.4278 | -94.1946
2.4315 | -94.1956
2.4352 | -94.2629
2.439 | -94.4098
2.4427 | -94.6038
2.4464 | -94.7888
2.4502 | -94.9014
2.4539 | -94.9286
2.4576 | -94.9301
2.4614 | -94.9622
2.4651 | -95.0542
2.4688 | -95.1514
2.4726 | -95.1891
2.4763 | -95.1169
2.48 | -94.923
2.4838 | -94.6762
2.4875 | -94.4742
2.4912 | -94.4043
2.495 | -94.5027
2.4987 | -94.7152
2.5025 | -94.9407
2.5062 | -95.074
2.5099 | -95.0652
2.5137 | -94.9446
2.5174 | -94.7852
2.5211 | -94.6704
2.5249 | -94.6349
2.5286 | -94.6486
2.5323 | -94.6561
2.5361 | -94.6415
2.5398 | -94.6133
2.5435 | -94.5962
2.5473 | -94.6235
2.551 | -94.723
2.5547 | -94.8727
2.5585 | -95.0275
2.5622 | -95.1416
2.5659 | -95.2128
2.5697 | -95.2418
2.5734 | -95.2303
2.5772 | -95.1618
2.5809 | -95.0521
2.5846 | -94.9502
2.5884 | -94.9106
2.5921 | -94.949
2.5958 | -95.0624
2.5996 | -95.2081
2.6033 | -95.3647
2.607 | -95.506
2.6108 | -95.6283
2.6145 | -95.716
2.6182 | -95.7781
2.622 | -95.8027
2.6257 | -95.7626
2.6294 | -95.6496
2.6332 | -95.4759
2.6369 | -95.2837
2.6406 | -95.1195
2.6444 | -94.9778
2.6481 | -94.8785
2.6519 | -94.8118
2.6556 | -94.7653
2.6593 | -94.7683
2.6631 | -94.8599
2.6668 | -95.0261
2.6705 | -95.2311
2.6743 | -95.4292
2.678 | -95.5712
2.6817 | -95.6378
2.6855 | -95.651
2.6892 | -95.6368
2.6929 | -95.6683
2.6967 | -95.6935
2.7004 | -95.7223
2.7041 | -95.7209
2.7079 | -95.6728
2.7116 | -95.5987
2.7153 | -95.5083
2.7191 | -95.4245
2.7228 | -95.3677
2.7266 | -95.3474
2.7303 | -95.3633
2.734 | -95.4131
2.7378 | -95.4737
2.7415 | -95.5024
2.7452 | -95.4862
2.749 | -95.4395
2.7527 | -95.3908
2.7564 | -95.3487
2.7602 | -95.326
2.7639 | -95.3171
2.7676 | -95.3147
2.7714 | -95.3174
2.7751 | -95.3412
2.7788 | -95.3963
2.7826 | -95.4836
2.7863 | -95.6098
2.79 | -95.7552
2.7938 | -95.8896
2.7975 | -95.9857
2.8013 | -96.0338
2.805 | -96.0532
2.8087 | -96.0441
2.8125 | -96.011
2.8162 | -95.9502
2.8199 | -95.8807
2.8237 | -95.8103
2.8274 | -95.7655
2.8311 | -95.7406
2.8349 | -95.7262
2.8386 | -95.7415
2.8423 | -95.8031
2.8461 | -95.9127
2.8498 | -96.0456
2.8535 | -96.162
2.8573 | -96.2254
2.861 | -96.251
2.8647 | -96.2519
2.8685 | -96.2191
2.8722 | -96.1701
2.876 | -96.1117
2.8797 | -96.0746
2.8834 | -96.0823
2.8872 | -96.1268
2.8909 | -96.206
2.8946 | -96.2961
2.8984 | -96.3653
2.9021 | -96.3969
2.9058 | -96.4
2.9096 | -96.387
2.9133 | -96.3852
2.917 | -96.4231
2.9208 | -96.4979
2.9245 | -96.5867
2.9282 | -96.6629
2.932 | -96.6862
2.9357 | -96.6471
2.9394 | -96.5308
2.9432 | -96.3574
2.9469 | -96.1822
2.9507 | -96.0502
2.9544 | -95.9887
2.9581 | -95.9977
2.9619 | -96.0421
2.9656 | -96.0679
2.9693 | -96.0454
2.9731 | -95.9914
2.9768 | -95.9197
2.9805 | -95.8441
2.9843 | -95.7719
2.988 | -95.7261
2.9917 | -95.7449
2.9955 | -95.8287
2.9992 | -95.9828
3.0029 | -96.1963
3.0067 | -96.426
3.0104 | -96.642
3.0141 | -96.8278
3.0179 | -96.9783
3.0216 | -97.0839
3.0254 | -97.1516
3.0291 | -97.1726
3.0328 | -97.1472
3.0366 | -97.0701
3.0403 | -96.9628
3.044 | -96.8537
3.0478 | -96.7375
3.0515 | -96.631
3.0552 | -96.5613
3.059 | -96.5207
3.0627 | -96.5125
3.0664 | -96.521
3.0702 | -96.5381
3.0739 | -96.5384
3.0776 | -96.4945
3.0814 | -96.4055
3.0851 | -96.3019
3.0888 | -96.216
3.0926 | -96.2052
3.0963 | -96.2814
3.1001 | -96.4154
3.1038 | -96.5695
3.1075 | -96.7009
3.1113 | -96.7862
3.115 | -96.8239
3.1187 | -96.7892
3.1225 | -96.6892
3.1262 | -96.5645
3.1299 | -96.4326
3.1337 | -96.3222
3.1374 | -96.2439
3.1411 | -96.191
3.1449 | -96.1662
3.1486 | -96.1688
3.1523 | -96.1945
3.1561 | -96.2528
3.1598 | -96.3358
3.1635 | -96.4349
3.1673 | -96.5509
3.171 | -96.6656
3.1748 | -96.7568
3.1785 | -96.8322
3.1822 | -96.8843
3.186 | -96.9253
3.1897 | -96.9626
3.1934 | -96.9821
3.1972 | -96.9803
3.2009 | -96.9642
3.2046 | -96.9389
3.2084 | -96.9175
3.2121 | -96.9189
3.2158 | -96.943
3.2196 | -96.9678
3.2233 | -96.9781
3.227 | -96.9588
3.2308 | -96.9122
3.2345 | -96.8883
3.2382 | -96.8667
3.242 | -96.865
3.2457 | -96.8872
3.2495 | -96.9222
3.2532 | -96.9695
3.2569 | -97.0281
3.2607 | -97.098
3.2644 | -97.1699
3.2681 | -97.245
3.2719 | -97.3234
3.2756 | -97.4
3.2793 | -97.4772
3.2831 | -97.5585
3.2868 | -97.6519
3.2905 | -97.7595
3.2943 | -97.8687
3.298 | -97.9517
3.3017 | -97.9998
3.3055 | -97.9969
3.3092 | -97.9288
3.3129 | -97.8001
3.3167 | -97.6364
3.3204 | -97.4706
3.3242 | -97.3372
3.3279 | -97.2756
3.3316 | -97.2934
3.3354 | -97.3612
3.3391 | -97.4457
3.3428 | -97.5246
3.3466 | -97.5564
3.3503 | -97.5503
3.354 | -97.5142
3.3578 | -97.4612
3.3615 | -97.4137
3.3652 | -97.3878
3.369 | -97.3996
3.3727 | -97.4497
3.3764 | -97.521
3.3802 | -97.5958
3.3839 | -97.6672
3.3876 | -97.7299
3.3914 | -97.7909
3.3951 | -97.8473
3.3989 | -97.8887
3.4026 | -97.9316
3.4063 | -97.9558
3.4101 | -97.9667
3.4138 | -97.971
3.4175 | -97.9719
3.4213 | -97.979
3.425 | -97.9828
3.4287 | -97.9732
3.4325 | -97.9435
3.4362 | -97.8898
3.4399 | -97.8381
3.4437 | -97.8005
3.4474 | -97.7992
3.4511 | -97.8474
3.4549 | -97.9423
3.4586 | -98.0739
3.4623 | -98.2288
3.4661 | -98.3791
3.4698 | -98.4972
3.4736 | -98.5569
3.4773 | -98.5428
3.481 | -98.4509
3.4848 | -98.3039
3.4885 | -98.1392
3.4922 | -97.9925
3.496 | -97.88
3.4997 | -97.8196
3.5034 | -97.797
3.5072 | -97.7948
3.5109 | -97.8215
3.5146 | -97.8647
3.5184 | -97.9189
3.5221 | -97.9718
3.5258 | -98.0128
3.5296 | -98.0243
3.5333 | -97.9895
3.537 | -97.921
3.5408 | -97.8458
3.5445 | -97.7873
3.5483 | -97.7512
3.552 | -97.7444
3.5557 | -97.7674
3.5595 | -97.8007
3.5632 | -97.818
3.5669 | -97.807
3.5707 | -97.7851
3.5744 | -97.7594
3.5781 | -97.7362
3.5819 | -97.7223
3.5856 | -97.7347
3.5893 | -97.7682
3.5931 | -97.807
3.5968 | -97.8335
3.6005 | -97.8571
3.6043 | -97.8813
3.608 | -97.8988
3.6117 | -97.9115
3.6155 | -97.925
3.6192 | -97.9457
3.623 | -97.9657
3.6267 | -97.9827
3.6304 | -98.0108
3.6342 | -98.059
3.6379 | -98.1183
3.6416 | -98.18
3.6454 | -98.2502
3.6491 | -98.3112
3.6528 | -98.3423
3.6566 | -98.3152
3.6603 | -98.2324
3.664 | -98.1056
3.6678 | -97.958
3.6715 | -97.8389
3.6752 | -97.7768
3.679 | -97.7878
3.6827 | -97.8783
3.6864 | -98.0247
3.6902 | -98.1972
3.6939 | -98.3577
3.6977 | -98.4752
3.7014 | -98.5178
3.7051 | -98.4952
3.7089 | -98.4281
3.7126 | -98.3479
3.7163 | -98.2873
3.7201 | -98.2722
3.7238 | -98.3173
3.7275 | -98.4048
3.7313 | -98.5259
3.735 | -98.6474
3.7387 | -98.7391
3.7425 | -98.7931
3.7462 | -98.8161
3.7499 | -98.8203
3.7537 | -98.8046
3.7574 | -98.7807
3.7611 | -98.7412
3.7649 | -98.6873
3.7686 | -98.6416
3.7724 | -98.6054
3.7761 | -98.5857
3.7798 | -98.5829
3.7836 | -98.5842
3.7873 | -98.5963
3.791 | -98.6113
3.7948 | -98.6203
3.7985 | -98.6172
3.8022 | -98.6144
3.806 | -98.6192
3.8097 | -98.6444
3.8134 | -98.6928
3.8172 | -98.7578
3.8209 | -98.8358
3.8246 | -98.9123
3.8284 | -98.9707
3.8321 | -99.0099
3.8358 | -99.023
3.8396 | -99.0126
3.8433 | -98.9992
3.8471 | -98.9861
3.8508 | -98.958
3.8545 | -98.9173
3.8583 | -98.8684
3.862 | -98.8008
3.8657 | -98.7142
3.8695 | -98.627
3.8732 | -98.5525
3.8769 | -98.4964
3.8807 | -98.4581
3.8844 | -98.4306
3.8881 | -98.4111
3.8919 | -98.4001
3.8956 | -98.3963
3.8993 | -98.4094
3.9031 | -98.448
3.9068 | -98.5118
3.9105 | -98.6133
3.9143 | -98.7456
3.918 | -98.8842
3.9218 | -99.0174
3.9255 | -99.1348
3.9292 | -99.1937
3.933 | -99.1891
3.9367 | -99.138
3.9404 | -99.0419
3.9442 | -98.9079
3.9479 | -98.7532
3.9516 | -98.6054
3.9554 | -98.48
3.9591 | -98.3992
3.9628 | -98.3625
3.9666 | -98.3763
3.9703 | -98.4299
3.974 | -98.5122
3.9778 | -98.6184
3.9815 | -98.7414
3.9852 | -98.8645
3.989 | -98.9811
3.9927 | -99.1096
3.9965 | -99.2208
4.0002 | -99.3166
4.0039 | -99.423
4.0077 | -99.5215
4.0114 | -99.591
4.0151 | -99.6078
4.0189 | -99.5762
4.0226 | -99.4973
4.0263 | -99.3856
4.0301 | -99.2531
4.0338 | -99.1244
4.0375 | -99.0031
4.0413 | -98.9005
4.045 | -98.8263
4.0487 | -98.7783
4.0525 | -98.746
4.0562 | -98.7183
4.0599 | -98.7038
4.0637 | -98.6922
4.0674 | -98.6698
4.0712 | -98.6561
4.0749 | -98.6564
4.0786 | -98.6715
4.0824 | -98.7063
4.0861 | -98.7609
4.0898 | -98.8396
4.0936 | -98.9259
4.0973 | -99.0143
4.101 | -99.0979
4.1048 | -99.1581
4.1085 | -99.1955
4.1122 | -99.2128
4.116 | -99.2146
4.1197 | -99.2062
4.1234 | -99.185
4.1272 | -99.1562
4.1309 | -99.13
4.1346 | -99.1051
4.1384 | -99.0845
4.1421 | -99.0665
4.1459 | -99.0316
4.1496 | -98.9827
4.1533 | -98.916
4.1571 | -98.8333
4.1608 | -98.7379
4.1645 | -98.6656
4.1683 | -98.6255
4.172 | -98.6102
4.1757 | -98.6294
4.1795 | -98.6745
4.1832 | -98.7266
4.1869 | -98.783
4.1907 | -98.8413
4.1944 | -98.9007
4.1981 | -98.9452
4.2019 | -98.9683
4.2056 | -98.9735
4.2093 | -98.9707
4.2131 | -98.9596
4.2168 | -98.9488
4.2206 | -98.9452
4.2243 | -98.949
4.228 | -98.955
4.2318 | -98.9745
4.2355 | -98.9991
4.2392 | -99.0248
4.243 | -99.054
4.2467 | -99.0959
4.2504 | -99.1493
4.2542 | -99.2099
4.2579 | -99.2818
4.2616 | -99.3729
4.2654 | -99.4701
4.2691 | -99.5572
4.2728 | -99.6214
4.2766 | -99.6575
4.2803 | -99.666
4.284 | -99.6588
4.2878 | -99.642
4.2915 | -99.6151
4.2953 | -99.5833
4.299 | -99.5526
4.3027 | -99.5096
4.3065 | -99.4534
4.3102 | -99.3993
4.3139 | -99.3404
4.3177 | -99.2887
4.3214 | -99.2526
4.3251 | -99.2234
4.3289 | -99.1864
4.3326 | -99.1558
4.3363 | -99.127
4.3401 | -99.1042
4.3438 | -99.0881
4.3475 | -99.0879
4.3513 | -99.1077
4.355 | -99.1479
4.3587 | -99.2012
4.3625 | -99.2701
4.3662 | -99.344
4.37 | -99.4151
4.3737 | -99.4807
4.3774 | -99.5416
4.3812 | -99.6046
4.3849 | -99.6821
4.3886 | -99.7583
4.3924 | -99.8264
4.3961 | -99.8964
4.3998 | -99.9735
4.4036 | -100.0325
4.4073 | -100.0814
4.411 | -100.1252
4.4148 | -100.1596
4.4185 | -100.1789
4.4222 | -100.1739
4.426 | -100.1392
4.4297 | -100.0912
4.4334 | -100.0219
4.4372 | -99.9434
4.4409 | -99.8664
4.4447 | -99.7917
4.4484 | -99.7189
4.4521 | -99.6477
4.4559 | -99.5861
4.4596 | -99.5473
4.4633 | -99.5343
4.4671 | -99.5495
4.4708 | -99.5915
4.4745 | -99.6656
4.4783 | -99.7572
4.482 | -99.8457
4.4857 | -99.9377
4.4895 | -100.0169
4.4932 | -100.0692
4.4969 | -100.0892
4.5007 | -100.0709
4.5044 | -100.0256
4.5081 | -99.9548
4.5119 | -99.8792
4.5156 | -99.8075
4.5194 | -99.7402
4.5231 | -99.6909
4.5268 | -99.6567
4.5306 | -99.6463
4.5343 | -99.6613
4.538 | -99.6907
4.5418 | -99.7399
4.5455 | -99.7979
4.5492 | -99.8562
4.553 | -99.8918
4.5567 | -99.8961
4.5604 | -99.8737
4.5642 | -99.8329
4.5679 | -99.7768
4.5716 | -99.7119
4.5754 | -99.6581
4.5791 | -99.621
4.5828 | -99.5938
4.5866 | -99.5949
4.5903 | -99.6209
4.5941 | -99.6673
4.5978 | -99.7294
4.6015 | -99.7919
4.6053 | -99.8455
4.609 | -99.8846
4.6127 | -99.9042
4.6165 | -99.9002
4.6202 | -99.8793
4.6239 | -99.8556
4.6277 | -99.8312
4.6314 | -99.813
4.6351 | -99.8014
4.6389 | -99.7976
4.6426 | -99.7992
4.6463 | -99.8033
4.6501 | -99.8181
4.6538 | -99.8374
4.6575 | -99.8552
4.6613 | -99.8852
4.665 | -99.9159
4.6688 | -99.9543
4.6725 | -100.0069
4.6762 | -100.0691
4.68 | -100.1352
4.6837 | -100.1993
4.6874 | -100.2738
4.6912 | -100.3478
4.6949 | -100.4135
4.6986 | -100.4661
4.7024 | -100.5052
4.7061 | -100.5324
4.7098 | -100.5473
4.7136 | -100.5488
4.7173 | -100.5274
4.721 | -100.4924
4.7248 | -100.456
4.7285 | -100.4151
4.7322 | -100.3733
4.736 | -100.3387
4.7397 | -100.3165
4.7435 | -100.3062
4.7472 | -100.3014
4.7509 | -100.3048
4.7547 | -100.3057
4.7584 | -100.2944
4.7621 | -100.2748
4.7659 | -100.2368
4.7696 | -100.1838
4.7733 | -100.1299
4.7771 | -100.0809
4.7808 | -100.0454
4.7845 | -100.0272
4.7883 | -100.0338
4.792 | -100.0602
4.7957 | -100.1033
4.7995 | -100.1572
4.8032 | -100.2138
4.8069 | -100.2736
4.8107 | -100.3287
4.8144 | -100.3708
4.8182 | -100.4031
4.8219 | -100.4133
4.8256 | -100.4041
4.8294 | -100.3715
4.8331 | -100.3279
4.8368 | -100.2911
4.8406 | -100.2584
4.8443 | -100.2337
4.848 | -100.2394
4.8518 | -100.2308
4.8555 | -100.2261
4.8592 | -100.23
4.863 | -100.2357
4.8667 | -100.235
4.8704 | -100.2242
4.8742 | -100.2005
4.8779 | -100.1717
4.8816 | -100.1304
4.8854 | -100.0847
4.8891 | -100.053
4.8929 | -100.0313
4.8966 | -100.0197
4.9003 | -100.0203
4.9041 | -100.0392
4.9078 | -100.068
4.9115 | -100.1158
4.9153 | -100.1848
4.919 | -100.2622
4.9227 | -100.3424
4.9265 | -100.4266
4.9302 | -100.5079
4.9339 | -100.5709
4.9377 | -100.6205
4.9414 | -100.6553
4.9451 | -100.673
4.9489 | -100.675
4.9526 | -100.6661
4.9563 | -100.637
4.9601 | -100.5924
4.9638 | -100.5266
4.9676 | -100.4476
4.9713 | -100.3685
4.975 | -100.3016
4.9788 | -100.2591
4.9825 | -100.2353
4.9862 | -100.2281
4.99 | -100.2416
4.9937 | -100.2751
4.9974 | -100.3304
5.0012 | -100.4044
5.0049 | -100.4793
5.0086 | -100.5603
5.0124 | -100.6385
5.0161 | -100.7032
5.0198 | -100.7573
5.0236 | -100.8043
5.0273 | -100.8389
5.031 | -100.8565
5.0348 | -100.8611
5.0385 | -100.8633
5.0423 | -100.867
5.046 | -100.8755
5.0497 | -100.8943
5.0535 | -100.9196
5.0572 | -100.946
5.0609 | -100.9867
5.0647 | -101.0347
5.0684 | -101.0781
5.0721 | -101.1229
5.0759 | -101.1608
5.0796 | -101.1906
5.0833 | -101.2127
5.0871 | -101.2265
5.0908 | -101.244
5.0945 | -101.2476
5.0983 | -101.2371
5.102 | -101.2094
5.1057 | -101.1714
5.1095 | -101.1239
5.1132 | -101.0647
5.117 | -100.9979
5.1207 | -100.9246
5.1244 | -100.8448
5.1282 | -100.7627
5.1319 | -100.6884
5.1356 | -100.6288
5.1394 | -100.5805
5.1431 | -100.557
5.1468 | -100.5539
5.1506 | -100.5703
5.1543 | -100.6029
5.158 | -100.6438
5.1618 | -100.695
5.1655 | -100.7341
5.1692 | -100.7669
5.173 | -100.7968
5.1767 | -100.8195
5.1804 | -100.8428
5.1842 | -100.8549
5.1879 | -100.8588
5.1917 | -100.8531
5.1954 | -100.8414
5.1991 | -100.8249
5.2029 | -100.8084
5.2066 | -100.7951
5.2103 | -100.7832
5.2141 | -100.7716
5.2178 | -100.7608
5.2215 | -100.739
5.2253 | -100.704
5.229 | -100.6599
5.2327 | -100.6131
5.2365 | -100.5729
5.2402 | -100.5386
5.2439 | -100.5216
5.2477 | -100.5144
5.2514 | -100.5165
5.2551 | -100.5256
5.2589 | -100.5376
5.2626 | -100.5545
5.2664 | -100.5721
5.2701 | -100.5941
5.2738 | -100.6124
5.2776 | -100.6298
5.2813 | -100.6465
5.285 | -100.6563
5.2888 | -100.6579
5.2925 | -100.6469
5.2962 | -100.6307
5.3 | -100.6154
5.3037 | -100.5993
5.3074 | -100.5828
5.3112 | -100.5672
5.3149 | -100.5542
5.3186 | -100.5546
5.3224 | -100.566
5.3261 | -100.5861
5.3298 | -100.6179
5.3336 | -100.6578
5.3373 | -100.708
5.3411 | -100.7747
5.3448 | -100.8456
5.3485 | -100.9168
5.3523 | -100.989
5.356 | -101.0649
5.3597 | -101.1378
5.3635 | -101.2047
5.3672 | -101.2648
5.3709 | -101.3206
5.3747 | -101.3623
5.3784 | -101.3766
5.3821 | -101.3709
5.3859 | -101.3191
5.3896 | -101.2854
5.3933 | -101.2478
5.3971 | -101.1983
5.4008 | -101.1398
5.4045 | -101.0853
5.4083 | -101.0393
5.412 | -100.9947
5.4158 | -100.9555
5.4195 | -100.926
5.4232 | -100.9034
5.427 | -100.8978
5.4307 | -100.9116
5.4344 | -100.935
5.4382 | -100.9663
5.4419 | -101.0083
5.4456 | -101.0577
5.4494 | -101.0985
5.4531 | -101.1382
5.4568 | -101.1775
5.4606 | -101.2169
5.4643 | -101.2503
5.468 | -101.2771
5.4718 | -101.2932
5.4755 | -101.3068
5.4792 | -101.2991
5.483 | -101.2806
5.4867 | -101.2533
5.4905 | -101.2122
5.4942 | -101.1626
5.4979 | -101.1137
5.5017 | -101.0696
5.5054 | -101.0377
5.5091 | -101.0261
5.5129 | -101.0356
5.5166 | -101.0586
5.5203 | -101.1089
5.5241 | -101.1776
5.5278 | -101.2543
5.5315 | -101.3353
5.5353 | -101.4137
5.539 | -101.471
5.5427 | -101.5071
5.5465 | -101.5303
5.5502 | -101.5351
5.5539 | -101.5217
5.5577 | -101.5006
5.5614 | -101.475
5.5652 | -101.4574
5.5689 | -101.4417
5.5726 | -101.4256
5.5764 | -101.4107
5.5801 | -101.398
5.5838 | -101.3876
5.5876 | -101.3736
5.5913 | -101.3573
5.595 | -101.3425
5.5988 | -101.3236
5.6025 | -101.3077
5.6062 | -101.2966
5.61 | -101.2874
5.6137 | -101.2843
5.6174 | -101.2874
5.6212 | -101.2986
5.6249 | -101.3128
5.6286 | -101.3304
5.6324 | -101.3413
5.6361 | -101.3548
5.6399 | -101.365
5.6436 | -101.3617
5.6473 | -101.3524
5.6511 | -101.3349
5.6548 | -101.3139
5.6585 | -101.291
5.6623 | -101.2764
5.666 | -101.2646
5.6697 | -101.2621
5.6735 | -101.2689
5.6772 | -101.2756
5.6809 | -101.2822
5.6847 | -101.2998
5.6884 | -101.3132
5.6921 | -101.3174
5.6959 | -101.323
5.6996 | -101.3298
5.7033 | -101.3341
5.7071 | -101.3305
5.7108 | -101.321
5.7146 | -101.3016
5.7183 | -101.2792
5.722 | -101.2492
5.7258 | -101.2273
5.7295 | -101.2192
5.7332 | -101.2159
5.737 | -101.222
5.7407 | -101.2354
5.7444 | -101.2565
5.7482 | -101.2831
5.7519 | -101.3198
5.7556 | -101.362
5.7594 | -101.4163
5.7631 | -101.4801
5.7668 | -101.547
5.7706 | -101.6143
5.7743 | -101.6854
5.778 | -101.749
5.7818 | -101.7997
5.7855 | -101.8352
5.7893 | -101.8597
5.793 | -101.869
5.7967 | -101.8692
5.8005 | -101.8623
5.8042 | -101.8479
5.8079 | -101.8395
5.8117 | -101.8304
5.8154 | -101.8217
5.8191 | -101.8199
5.8229 | -101.8253
5.8266 | -101.8282
5.8303 | -101.8346
5.8341 | -101.8477
5.8378 | -101.8638
5.8415 | -101.8841
5.8453 | -101.8969
5.849 | -101.9049
5.8527 | -101.915
5.8565 | -101.9111
5.8602 | -101.8975
5.864 | -101.8799
5.8677 | -101.857
5.8714 | -101.8348
5.8752 | -101.8202
5.8789 | -101.8234
5.8826 | -101.8351
5.8864 | -101.864
5.8901 | -101.9153
5.8938 | -101.9762
5.8976 | -102.0482
5.9013 | -102.1286
5.905 | -102.2119
5.9088 | -102.2791
5.9125 | -102.3304
5.9162 | -102.3589
5.92 | -102.3564
5.9237 | -102.2916
5.9274 | -102.2603
5.9312 | -102.216
5.9349 | -102.1613
5.9387 | -102.1065
5.9424 | -102.047
5.9461 | -101.9948
5.9499 | -101.9514
5.9536 | -101.9181
5.9573 | -101.8947
5.9611 | -101.8788
5.9648 | -101.8633
5.9685 | -101.8555
5.9723 | -101.854
5.976 | -101.8542
5.9797 | -101.8518
5.9835 | -101.853
5.9872 | -101.8607
5.9909 | -101.8787
5.9947 | -101.9013
5.9984 | -101.935
6.0021 | -101.9721
6.0059 | -102.0074
6.0096 | -102.043
6.0134 | -102.0755
6.0171 | -102.0987
6.0208 | -102.1175
6.0246 | -102.1319
6.0283 | -102.149
6.032 | -102.1573
6.0358 | -102.1683
6.0395 | -102.1852
6.0432 | -102.2025
6.047 | -102.2206
6.0507 | -102.2362
6.0544 | -102.244
6.0582 | -102.2526
6.0619 | -102.245
6.0656 | -102.2236
6.0694 | -102.192
6.0731 | -102.1566
6.0768 | -102.1124
6.0806 | -102.0635
6.0843 | -102.0104
6.0881 | -101.9636
6.0918 | -101.9209
6.0955 | -101.8892
6.0993 | -101.8595
6.103 | -101.837
6.1067 | -101.8195
6.1105 | -101.8046
6.1142 | -101.7888
6.1179 | -101.7732
6.1217 | -101.7552
6.1254 | -101.7415
6.1291 | -101.7282
6.1329 | -101.7091
6.1366 | -101.6926
6.1403 | -101.6805
6.1441 | -101.6815
6.1478 | -101.6917
6.1515 | -101.6993
6.1553 | -101.7097
6.159 | -101.7339
6.1628 | -101.7647
6.1665 | -101.7916
6.1702 | -101.8206
6.174 | -101.8548
6.1777 | -101.8931
6.1814 | -101.9226
6.1852 | -101.9538
6.1889 | -101.9864
6.1926 | -102.0225
6.1964 | -102.0503
6.2001 | -102.0731
6.2038 | -102.0909
6.2076 | -102.1082
6.2113 | -102.119
6.215 | -102.122
6.2188 | -102.1196
6.2225 | -102.1121
6.2262 | -102.1068
6.23 | -102.1022
6.2337 | -102.0998
6.2375 | -102.0867
6.2412 | -102.069
6.2449 | -102.0562
6.2487 | -102.0438
6.2524 | -102.0244
6.2561 | -102.0037
6.2599 | -101.9921
6.2636 | -101.9864
6.2673 | -101.9855
6.2711 | -101.9942
6.2748 | -102.0143
6.2785 | -102.0343
6.2823 | -102.055
6.286 | -102.0855
6.2897 | -102.1263
6.2935 | -102.1643
6.2972 | -102.1963
6.3009 | -102.2278
6.3047 | -102.2594
6.3084 | -102.2929
6.3122 | -102.3223
6.3159 | -102.3433
6.3196 | -102.3581
6.3234 | -102.3573
6.3271 | -102.3377
6.3308 | -102.3148
6.3346 | -102.2843
6.3383 | -102.2485
6.342 | -102.2013
6.3458 | -102.1475
6.3495 | -102.0997
6.3532 | -102.0491
6.357 | -102.0092
6.3607 | -101.9771
6.3644 | -101.9541
6.3682 | -101.9469
6.3719 | -101.9539
6.3756 | -101.9688
6.3794 | -101.9885
6.3831 | -102.0152
6.3869 | -102.0494
6.3906 | -102.0838
6.3943 | -102.1161
6.3981 | -102.1443
6.4018 | -102.1685
6.4055 | -102.1915
6.4093 | -102.2171
6.413 | -102.2406
6.4167 | -102.2667
6.4205 | -102.2971
6.4242 | -102.3268
6.4279 | -102.3593
6.4317 | -102.3905
6.4354 | -102.4315
6.4391 | -102.4706
6.4429 | -102.4964
6.4466 | -102.5258
6.4503 | -102.5415
6.4541 | -102.5482
6.4578 | -102.5524
6.4616 | -102.5462
6.4653 | -102.5188
6.469 | -102.5053
6.4728 | -102.4873
6.4765 | -102.4664
6.4802 | -102.4442
6.484 | -102.4199
6.4877 | -102.3992
6.4914 | -102.3753
6.4952 | -102.3588
6.4989 | -102.3472
6.5026 | -102.3409
6.5064 | -102.3406
6.5101 | -102.3475
6.5138 | -102.3662
6.5176 | -102.3953
6.5213 | -102.4292
6.525 | -102.4768
6.5288 | -102.5333
6.5325 | -102.5966
6.5363 | -102.668
6.54 | -102.7465
6.5437 | -102.8224
6.5475 | -102.8962
6.5512 | -102.9665
6.5549 | -103.0293
6.5587 | -103.0835
6.5624 | -103.124
6.5661 | -103.1486
6.5699 | -103.1563
6.5736 | -103.1429
6.5773 | -103.1137
6.5811 | -103.0784
6.5848 | -103.034
6.5885 | -102.9716
6.5923 | -102.9031
6.596 | -102.8256
6.5997 | -102.7483
6.6035 | -102.6726
6.6072 | -102.5989
6.611 | -102.5284
6.6147 | -102.4667
6.6184 | -102.4126
6.6222 | -102.364
6.6259 | -102.32
6.6296 | -102.2866
6.6334 | -102.2677
6.6371 | -102.2565
6.6408 | -102.2515
6.6446 | -102.2464
6.6483 | -102.2409
6.652 | -102.2348
6.6558 | -102.2299
6.6595 | -102.2266
6.6632 | -102.2268
6.667 | -102.2274
6.6707 | -102.2353
6.6744 | -102.2566
6.6782 | -102.2814
6.6819 | -102.3105
6.6857 | -102.348
6.6894 | -102.397
6.6931 | -102.4492
6.6969 | -102.5027
6.7006 | -102.5567
6.7043 | -102.6118
6.7081 | -102.6638
6.7118 | -102.7116
6.7155 | -102.7533
6.7193 | -102.7824
6.723 | -102.8055
6.7267 | -102.8194
6.7305 | -102.8114
6.7342 | -102.7882
6.7379 | -102.7622
6.7417 | -102.7302
6.7454 | -102.6986
6.7491 | -102.6635
6.7529 | -102.6261
6.7566 | -102.5856
6.7604 | -102.5573
6.7641 | -102.5338
6.7678 | -102.5162
6.7716 | -102.5026
6.7753 | -102.4943
6.779 | -102.4979
6.7828 | -102.5027
6.7865 | -102.5128
6.7902 | -102.5268
6.794 | -102.5464
6.7977 | -102.5703
6.8014 | -102.5935
6.8052 | -102.6165
6.8089 | -102.6408
6.8126 | -102.6603
6.8164 | -102.6812
6.8201 | -102.7082
6.8238 | -102.7321
6.8276 | -102.7535
6.8313 | -102.7835
6.8351 | -102.8111
6.8388 | -102.8298
6.8425 | -102.8491
6.8463 | -102.8669
6.85 | -102.8873
6.8537 | -102.9042
6.8575 | -102.9144
6.8612 | -102.9088
6.8649 | -102.903
6.8687 | -102.8985
6.8724 | -102.8965
6.8761 | -102.8915
6.8799 | -102.8853
6.8836 | -102.8884
6.8873 | -102.8937
6.8911 | -102.9019
6.8948 | -102.9057
6.8985 | -102.9148
6.9023 | -102.931
6.906 | -102.9466
6.9098 | -102.9678
6.9135 | -102.9999
6.9172 | -103.0276
6.921 | -103.0582
6.9247 | -103.0923
6.9284 | -103.1179
6.9322 | -103.1399
6.9359 | -103.1558
6.9396 | -103.1734
6.9434 | -103.1767
6.9471 | -103.1735
6.9508 | -103.1656
6.9546 | -103.1513
6.9583 | -103.1414
6.962 | -103.134
6.9658 | -103.1214
6.9695 | -103.1077
6.9732 | -103.0941
6.977 | -103.0849
6.9807 | -103.0793
6.9845 | -103.0723
6.9882 | -103.0698
6.9919 | -103.0768
6.9957 | -103.0989
6.9994 | -103.1202
7.0031 | -103.1505
7.0069 | -103.1747
7.0106 | -103.2037
7.0143 | -103.2291
7.0181 | -103.2523
7.0218 | -103.2733
7.0255 | -103.2866
7.0293 | -103.2933
7.033 | -103.3059
7.0367 | -103.3064
7.0405 | -103.2995
7.0442 | -103.2872
7.0479 | -103.2711
7.0517 | -103.2457
7.0554 | -103.2181
7.0592 | -103.1903
7.0629 | -103.1558
7.0666 | -103.1203
7.0704 | -103.0879
7.0741 | -103.0554
7.0778 | -103.0251
7.0816 | -103.0044
7.0853 | -102.9849
7.089 | -102.9721
7.0928 | -102.9675
7.0965 | -102.9658
7.1002 | -102.9713
7.104 | -102.9816
7.1077 | -102.9938
7.1114 | -103.0061
7.1152 | -103.0185
7.1189 | -103.0316
7.1226 | -103.0429
7.1264 | -103.0471
7.1301 | -103.0505
7.1339 | -103.052
7.1376 | -103.0506
7.1413 | -103.0313
7.1451 | -103.0135
7.1488 | -102.9907
7.1525 | -102.9573
7.1563 | -102.9244
7.16 | -102.8898
7.1637 | -102.8495
7.1675 | -102.8109
7.1712 | -102.7799
7.1749 | -102.7556
7.1787 | -102.7385
7.1824 | -102.7349
7.1861 | -102.7409
7.1899 | -102.7539
7.1936 | -102.7838
7.1973 | -102.8208
7.2011 | -102.8649
7.2048 | -102.916
7.2086 | -102.9683
7.2123 | -103.029
7.216 | -103.0894
7.2198 | -103.1343
7.2235 | -103.1636
7.2272 | -103.1874
7.231 | -103.1985
7.2347 | -103.197
7.2384 | -103.1808
7.2422 | -103.1531
7.2459 | -103.113
7.2496 | -103.0737
7.2534 | -103.0315
7.2571 | -102.9912
7.2608 | -102.9516
7.2646 | -102.9096
7.2683 | -102.8729
7.272 | -102.8443
7.2758 | -102.8189
7.2795 | -102.7898
7.2833 | -102.769
7.287 | -102.7477
7.2907 | -102.7259
7.2945 | -102.7106
7.2982 | -102.6966
7.3019 | -102.6779
7.3057 | -102.6635
7.3094 | -102.6563
7.3131 | -102.654
7.3169 | -102.6582
7.3206 | -102.6656
7.3243 | -102.6809
7.3281 | -102.7125
7.3318 | -102.7475
7.3355 | -102.7824
7.3393 | -102.8278
7.343 | -102.876
7.3467 | -102.927
7.3505 | -102.978
7.3542 | -103.03
7.358 | -103.0785
7.3617 | -103.1189
7.3654 | -103.1524
7.3692 | -103.1831
7.3729 | -103.2122
7.3766 | -103.2339
7.3804 | -103.251
7.3841 | -103.2675
7.3878 | -103.2829
7.3916 | -103.2881
7.3953 | -103.29
7.399 | -103.2847
7.4028 | -103.2762
7.4065 | -103.2665
7.4102 | -103.2611
7.414 | -103.256
7.4177 | -103.2448
7.4214 | -103.2408
7.4252 | -103.2464
7.4289 | -103.2598
7.4327 | -103.2814
7.4364 | -103.3066
7.4401 | -103.3382
7.4439 | -103.3706
7.4476 | -103.4044
7.4513 | -103.4424
7.4551 | -103.4774
7.4588 | -103.5135
7.4625 | -103.546
7.4663 | -103.5747
7.47 | -103.5961
7.4737 | -103.6082
7.4775 | -103.6093
7.4812 | -103.5963
7.4849 | -103.568
7.4887 | -103.5351
7.4924 | -103.4957
7.4961 | -103.455
7.4999 | -103.412
7.5036 | -103.3722
7.5074 | -103.3392
7.5111 | -103.3042
7.5148 | -103.275
7.5186 | -103.254
7.5223 | -103.2407
7.526 | -103.2379
7.5298 | -103.2368
7.5335 | -103.2349
7.5372 | -103.2359
7.541 | -103.2559
7.5447 | -103.2651
7.5484 | -103.279
7.5522 | -103.2964
7.5559 | -103.3161
7.5596 | -103.3314
7.5634 | -103.3485
7.5671 | -103.3653
7.5708 | -103.3768
7.5746 | -103.3834
7.5783 | -103.3874
7.5821 | -103.3885
7.5858 | -103.3895
7.5895 | -103.3923
7.5933 | -103.3967
7.597 | -103.402
7.6007 | -103.4024
7.6045 | -103.4051
7.6082 | -103.4098
7.6119 | -103.4175
7.6157 | -103.4203
7.6194 | -103.4244
7.6231 | -103.4362
7.6269 | -103.4501
7.6306 | -103.4632
7.6343 | -103.4795
7.6381 | -103.4983
7.6418 | -103.5188
7.6455 | -103.5399
7.6493 | -103.5593
7.653 | -103.5695
7.6568 | -103.5819
7.6605 | -103.5857
7.6642 | -103.5866
7.668 | -103.5851
7.6717 | -103.5803
7.6754 | -103.5803
7.6792 | -103.583
7.6829 | -103.5864
7.6866 | -103.5951
7.6904 | -103.6084
7.6941 | -103.6254
7.6978 | -103.6502
7.7016 | -103.6756
7.7053 | -103.7032
7.709 | -103.7329
7.7128 | -103.7657
7.7165 | -103.7986
7.7202 | -103.8299
7.724 | -103.8597
7.7277 | -103.8808
7.7315 | -103.8976
7.7352 | -103.8995
7.7389 | -103.885
7.7427 | -103.8685
7.7464 | -103.8473
7.7501 | -103.8193
7.7539 | -103.7882
7.7576 | -103.7564
7.7613 | -103.7257
7.7651 | -103.6967
7.7688 | -103.6667
7.7725 | -103.6361
7.7763 | -103.6126
7.78 | -103.5944
7.7837 | -103.5828
7.7875 | -103.5753
7.7912 | -103.5671
7.7949 | -103.5616
7.7987 | -103.5636
7.8024 | -103.5637
7.8062 | -103.5612
7.8099 | -103.5591
7.8136 | -103.5562
7.8174 | -103.5564
7.8211 | -103.5547
7.8248 | -103.5512
7.8286 | -103.5509
7.8323 | -103.5522
7.836 | -103.5565
7.8398 | -103.561
7.8435 | -103.569
7.8472 | -103.5787
7.851 | -103.5949
7.8547 | -103.6145
7.8584 | -103.6284
7.8622 | -103.6428
7.8659 | -103.6584
7.8696 | -103.6721
7.8734 | -103.6862
7.8771 | -103.6995
7.8809 | -103.7114
7.8846 | -103.7247
7.8883 | -103.7359
7.8921 | -103.7438
7.8958 | -103.7456
7.8995 | -103.7406
7.9033 | -103.7317
7.907 | -103.7291
7.9107 | -103.722
7.9145 | -103.7139
7.9182 | -103.71
7.9219 | -103.706
7.9257 | -103.696
7.9294 | -103.6868
7.9331 | -103.6781
7.9369 | -103.6704
7.9406 | -103.6656
7.9443 | -103.6589
7.9481 | -103.6518
7.9518 | -103.6465
7.9556 | -103.6456
7.9593 | -103.6478
7.963 | -103.6471
7.9668 | -103.6376
7.9705 | -103.6279
7.9742 | -103.6188
7.978 | -103.6062
7.9817 | -103.5925
7.9854 | -103.5811
7.9892 | -103.5695
7.9929 | -103.5575
7.9966 | -103.5454
8.0004 | -103.5352
8.0041 | -103.5267
8.0078 | -103.5191
8.0116 | -103.5167
8.0153 | -103.5233
8.019 | -103.5354
8.0228 | -103.5446
8.0265 | -103.5603
8.0303 | -103.5806
8.034 | -103.6015
8.0377 | -103.6225
8.0415 | -103.6491
8.0452 | -103.6784
8.0489 | -103.7088
8.0527 | -103.736
8.0564 | -103.7594
8.0601 | -103.7788
8.0639 | -103.7934
8.0676 | -103.8
8.0713 | -103.801
8.0751 | -103.7995
8.0788 | -103.7997
8.0825 | -103.7902
8.0863 | -103.7827
8.09 | -103.7675
8.0937 | -103.7506
8.0975 | -103.7301
8.1012 | -103.7044
8.105 | -103.675
8.1087 | -103.6446
8.1124 | -103.6087
8.1162 | -103.5699
8.1199 | -103.5247
8.1236 | -103.4841
8.1274 | -103.4487
8.1311 | -103.4084
8.1348 | -103.3692
8.1386 | -103.3378
8.1423 | -103.3185
8.146 | -103.3078
8.1498 | -103.3079
8.1535 | -103.3169
8.1572 | -103.3385
8.161 | -103.3618
8.1647 | -103.3861
8.1684 | -103.415
8.1722 | -103.4516
8.1759 | -103.4904
8.1797 | -103.5254
8.1834 | -103.5638
8.1871 | -103.6048
8.1909 | -103.6435
8.1946 | -103.6776
8.1983 | -103.7073
8.2021 | -103.7416
8.2058 | -103.7789
8.2095 | -103.8178
8.2133 | -103.8526
8.217 | -103.8785
8.2207 | -103.9068
8.2245 | -103.9317
8.2282 | -103.9507
8.2319 | -103.9636
8.2357 | -103.9733
8.2394 | -103.9786
8.2431 | -103.9838
8.2469 | -103.9818
8.2506 | -103.9725
8.2544 | -103.9634
8.2581 | -103.9545
8.2618 | -103.9403
8.2656 | -103.9212
8.2693 | -103.9071
8.273 | -103.8889
8.2768 | -103.8771
8.2805 | -103.8666
8.2842 | -103.8621
8.288 | -103.8642
8.2917 | -103.8684
8.2954 | -103.8767
8.2992 | -103.8907
8.3029 | -103.9113
8.3066 | -103.9289
8.3104 | -103.9498
8.3141 | -103.9776
8.3178 | -104.0097
8.3216 | -104.0438
8.3253 | -104.0713
8.3291 | -104.0972
8.3328 | -104.1221
8.3365 | -104.1501
8.3403 | -104.1778
8.344 | -104.1984
8.3477 | -104.214
8.3515 | -104.226
8.3552 | -104.2346
8.3589 | -104.2424
8.3627 | -104.2515
8.3664 | -104.256
8.3701 | -104.2596
8.3739 | -104.2653
8.3776 | -104.2701
8.3813 | -104.2768
8.3851 | -104.2876
8.3888 | -104.2942
8.3925 | -104.2969
8.3963 | -104.3012
8.4 | -104.3033
8.4038 | -104.3083
8.4075 | -104.3114
8.4112 | -104.3102
8.415 | -104.3115
8.4187 | -104.3088
8.4224 | -104.3009
8.4262 | -104.2866
8.4299 | -104.2702
8.4336 | -104.2495
8.4374 | -104.2195
8.4411 | -104.1874
8.4448 | -104.1499
8.4486 | -104.112
8.4523 | -104.0756
8.456 | -104.0396
8.4598 | -104.0043
8.4635 | -103.9703
8.4672 | -103.9369
8.471 | -103.911
8.4747 | -103.8926
8.4785 | -103.8768
8.4822 | -103.8644
8.4859 | -103.8513
8.4897 | -103.8406
8.4934 | -103.8328
8.4971 | -103.8266
8.5009 | -103.8251
8.5046 | -103.8206
8.5083 | -103.8151
8.5121 | -103.8136
8.5158 | -103.8174
8.5195 | -103.8255
8.5233 | -103.8304
8.527 | -103.8351
8.5307 | -103.8415
8.5345 | -103.8476
8.5382 | -103.8595
8.5419 | -103.8773
8.5457 | -103.8974
8.5494 | -103.9191
8.5532 | -103.9342
8.5569 | -103.9472
8.5606 | -103.9644
8.5644 | -103.9815
8.5681 | -104.0011
8.5718 | -104.0195
8.5756 | -104.0294
8.5793 | -104.0364
8.583 | -104.0437
8.5868 | -104.0499
8.5905 | -104.0565
8.5942 | -104.0593
8.598 | -104.0579
8.6017 | -104.0651
8.6054 | -104.0738
8.6092 | -104.0872
8.6129 | -104.1062
8.6166 | -104.1206
8.6204 | -104.1392
8.6241 | -104.1588
8.6279 | -104.1757
8.6316 | -104.1941
8.6353 | -104.2103
8.6391 | -104.2221
8.6428 | -104.2279
8.6465 | -104.2336
8.6503 | -104.2397
8.654 | -104.24
8.6577 | -104.2417
8.6615 | -104.2438
8.6652 | -104.2407
8.6689 | -104.2385
8.6727 | -104.237
8.6764 | -104.2361
8.6801 | -104.2351
8.6839 | -104.2365
8.6876 | -104.2367
8.6913 | -104.2365
8.6951 | -104.2381
8.6988 | -104.2406
8.7026 | -104.2451
8.7063 | -104.253
8.71 | -104.2572
8.7138 | -104.2596
8.7175 | -104.2594
8.7212 | -104.26
8.725 | -104.2582
8.7287 | -104.2483
8.7324 | -104.2342
8.7362 | -104.2153
8.7399 | -104.1982
8.7436 | -104.1858
8.7474 | -104.1715
8.7511 | -104.151
8.7548 | -104.1294
8.7586 | -104.1149
8.7623 | -104.1009
8.766 | -104.0878
8.7698 | -104.0778
8.7735 | -104.0735
8.7773 | -104.0707
8.781 | -104.0738
8.7847 | -104.0855
8.7885 | -104.1
8.7922 | -104.1194
8.7959 | -104.1396
8.7997 | -104.1636
8.8034 | -104.1926
8.8071 | -104.2244
8.8109 | -104.262
8.8146 | -104.2957
8.8183 | -104.3272
8.8221 | -104.3576
8.8258 | -104.388
8.8295 | -104.4139
8.8333 | -104.4363
8.837 | -104.4552
8.8407 | -104.4743
8.8445 | -104.4934
8.8482 | -104.5115
8.852 | -104.5193
8.8557 | -104.5252
8.8594 | -104.5326
8.8632 | -104.5436
8.8669 | -104.5537
8.8706 | -104.5587
8.8744 | -104.5633
8.8781 | -104.57
8.8818 | -104.5809
8.8856 | -104.5974
8.8893 | -104.6135
8.893 | -104.6292
8.8968 | -104.6491
8.9005 | -104.6674
8.9042 | -104.6762
8.908 | -104.6874
8.9117 | -104.7001
8.9154 | -104.7061
8.9192 | -104.7106
8.9229 | -104.7123
8.9267 | -104.7091
8.9304 | -104.6984
8.9341 | -104.6852
8.9379 | -104.6687
8.9416 | -104.6456
8.9453 | -104.6253
8.9491 | -104.6032
8.9528 | -104.579
8.9565 | -104.552
8.9603 | -104.5282
8.964 | -104.5047
8.9677 | -104.4797
8.9715 | -104.4517
8.9752 | -104.4243
8.9789 | -104.4046
8.9827 | -104.3926
8.9864 | -104.3819
8.9901 | -104.3703
8.9939 | -104.3613
8.9976 | -104.3591
9.0014 | -104.3629
9.0051 | -104.372
9.0088 | -104.3856
9.0126 | -104.4054
9.0163 | -104.4283
9.02 | -104.4549
9.0238 | -104.4861
9.0275 | -104.5209
9.0312 | -104.5573
9.035 | -104.5963
9.0387 | -104.6305
9.0424 | -104.6564
9.0462 | -104.6827
9.0499 | -104.7073
9.0536 | -104.7238
9.0574 | -104.7374
9.0611 | -104.7439
9.0648 | -104.7415
9.0686 | -104.7342
9.0723 | -104.7217
9.0761 | -104.7032
9.0798 | -104.6829
9.0835 | -104.6626
9.0873 | -104.6424
9.091 | -104.6238
9.0947 | -104.6079
9.0985 | -104.5954
9.1022 | -104.5851
9.1059 | -104.5771
9.1097 | -104.572
9.1134 | -104.5769
9.1171 | -104.5872
9.1209 | -104.6004
9.1246 | -104.6155
9.1283 | -104.6337
9.1321 | -104.6467
9.1358 | -104.6599
9.1395 | -104.6747
9.1433 | -104.6803
9.147 | -104.6847
9.1508 | -104.6881
9.1545 | -104.6914
9.1582 | -104.6715
9.162 | -104.66
9.1657 | -104.6453
9.1694 | -104.6278
9.1732 | -104.6089
9.1769 | -104.59
9.1806 | -104.5664
9.1844 | -104.5406
9.1881 | -104.5176
9.1918 | -104.4948
9.1956 | -104.4765
9.1993 | -104.458
9.203 | -104.4392
9.2068 | -104.4273
9.2105 | -104.4175
9.2142 | -104.4144
9.218 | -104.4178
9.2217 | -104.43
9.2255 | -104.4487
9.2292 | -104.4727
9.2329 | -104.498
9.2367 | -104.5324
9.2404 | -104.5723
9.2441 | -104.6084
9.2479 | -104.6415
9.2516 | -104.6744
9.2553 | -104.7079
9.2591 | -104.7382
9.2628 | -104.7603
9.2665 | -104.7782
9.2703 | -104.7847
9.274 | -104.7891
9.2777 | -104.7933
9.2815 | -104.7922
9.2852 | -104.7826
9.2889 | -104.7657
9.2927 | -104.7476
9.2964 | -104.7288
9.3002 | -104.7068
9.3039 | -104.6836
9.3076 | -104.66
9.3114 | -104.6357
9.3151 | -104.6076
9.3188 | -104.5808
9.3226 | -104.5532
9.3263 | -104.5325
9.33 | -104.5097
9.3338 | -104.4853
9.3375 | -104.4679
9.3412 | -104.4558
9.345 | -104.443
9.3487 | -104.4329
9.3524 | -104.4246
9.3562 | -104.4171
9.3599 | -104.4078
9.3636 | -104.3993
9.3674 | -104.3954
9.3711 | -104.3947
9.3749 | -104.3933
9.3786 | -104.3914
9.3823 | -104.3938
9.3861 | -104.4048
9.3898 | -104.4208
9.3935 | -104.434
9.3973 | -104.4463
9.401 | -104.4675
9.4047 | -104.4952
9.4085 | -104.5212
9.4122 | -104.5354
9.4159 | -104.5477
9.4197 | -104.5552
9.4234 | -104.5634
9.4271 | -104.5728
9.4309 | -104.5744
9.4346 | -104.5694
9.4383 | -104.5605
9.4421 | -104.5465
9.4458 | -104.5246
9.4496 | -104.5023
9.4533 | -104.4794
9.457 | -104.4531
9.4608 | -104.4235
9.4645 | -104.3899
9.4682 | -104.3588
9.472 | -104.3319
9.4757 | -104.3048
9.4794 | -104.2782
9.4832 | -104.2577
9.4869 | -104.24
9.4906 | -104.2272
9.4944 | -104.2214
9.4981 | -104.2226
9.5018 | -104.2294
9.5056 | -104.2394
9.5093 | -104.2548
9.513 | -104.2766
9.5168 | -104.3033
9.5205 | -104.3327
9.5243 | -104.3592
9.528 | -104.3904
9.5317 | -104.4295
9.5355 | -104.4715
9.5392 | -104.507
9.5429 | -104.5411
9.5467 | -104.5779
9.5504 | -104.6219
9.5541 | -104.6643
9.5579 | -104.701
9.5616 | -104.7365
9.5653 | -104.7657
9.5691 | -104.7929
9.5728 | -104.8163
9.5765 | -104.831
9.5803 | -104.8457
9.584 | -104.8584
9.5877 | -104.8625
9.5915 | -104.8637
9.5952 | -104.8617
9.599 | -104.8564
9.6027 | -104.8464
9.6064 | -104.8364
9.6102 | -104.8207
9.6139 | -104.7981
9.6176 | -104.7781
9.6214 | -104.7607
9.6251 | -104.7418
9.6288 | -104.7249
9.6326 | -104.7054
9.6363 | -104.6891
9.64 | -104.6699
9.6438 | -104.6538
9.6475 | -104.6424
9.6512 | -104.6315
9.655 | -104.6239
9.6587 | -104.6145
9.6624 | -104.606
9.6662 | -104.5988
9.6699 | -104.5932
9.6737 | -104.584
9.6774 | -104.5769
9.6811 | -104.5728
9.6849 | -104.565
9.6886 | -104.5496
9.6923 | -104.5348
9.6961 | -104.5116
9.6998 | -104.4982
9.7035 | -104.4815
9.7073 | -104.4636
9.711 | -104.4471
9.7147 | -104.4323
9.7185 | -104.4213
9.7222 | -104.4078
9.7259 | -104.3997
9.7297 | -104.3977
9.7334 | -104.4001
9.7371 | -104.4027
9.7409 | -104.4108
9.7446 | -104.4226
9.7484 | -104.435
9.7521 | -104.4502
9.7558 | -104.4703
9.7596 | -104.4895
9.7633 | -104.5061
9.767 | -104.5251
9.7708 | -104.5463
9.7745 | -104.5656
9.7782 | -104.5811
9.782 | -104.6007
9.7857 | -104.619
9.7894 | -104.6348
9.7932 | -104.6495
9.7969 | -104.6596
9.8006 | -104.675
9.8044 | -104.6918
9.8081 | -104.6988
9.8118 | -104.7028
9.8156 | -104.7152
9.8193 | -104.7257
9.8231 | -104.7311
9.8268 | -104.7378
9.8305 | -104.7447
9.8343 | -104.7446
9.838 | -104.7433
9.8417 | -104.7458
9.8455 | -104.7444
9.8492 | -104.7407
9.8529 | -104.7345
9.8567 | -104.7257
9.8604 | -104.7208
9.8641 | -104.7104
9.8679 | -104.7001
9.8716 | -104.6927
9.8753 | -104.6873
9.8791 | -104.682
9.8828 | -104.6817
9.8865 | -104.6845
9.8903 | -104.6876
9.894 | -104.6944
9.8978 | -104.703
9.9015 | -104.7136
9.9052 | -104.7212
9.909 | -104.7261
9.9127 | -104.7328
9.9164 | -104.7392
9.9202 | -104.746
9.9239 | -104.7549
9.9276 | -104.7628
9.9314 | -104.7695
9.9351 | -104.7719
9.9388 | -104.7795
9.9426 | -104.7886
9.9463 | -104.7925
9.95 | -104.7966
9.9538 | -104.8049
9.9575 | -104.8084
9.9612 | -104.8068
9.965 | -104.8066
9.9687 | -104.8066
9.9725 | -104.8076
9.9762 | -104.8096
9.9799 | -104.816
9.9837 | -104.8232
9.9874 | -104.8262
9.9911 | -104.8334
9.9949 | -104.8446
9.9986 | -104.8594
10.0023 | -104.8774
10.0061 | -104.8931
10.0098 | -104.9157
10.0135 | -104.9374
10.0173 | -104.9565
10.021 | -104.9746
10.0247 | -104.9923
10.0285 | -105.009
10.0322 | -105.0241
10.0359 | -105.0378
10.0397 | -105.0483
10.0434 | -105.0576
10.0472 | -105.0619
10.0509 | -105.0667
10.0546 | -105.0693
10.0584 | -105.0656
10.0621 | -105.0522
10.0658 | -105.0389
10.0696 | -105.0242
10.0733 | -105.0069
10.077 | -104.9859
10.0808 | -104.9676
10.0845 | -104.9437
10.0882 | -104.9209
10.092 | -104.8962
10.0957 | -104.8657
10.0994 | -104.8398
10.1032 | -104.8185
10.1069 | -104.7985
10.1106 | -104.7818
10.1144 | -104.7693
10.1181 | -104.7634
10.1219 | -104.7623
10.1256 | -104.7642
10.1293 | -104.7725
10.1331 | -104.7862
10.1368 | -104.8033
10.1405 | -104.8233
10.1443 | -104.848
10.148 | -104.8728
10.1517 | -104.9015
10.1555 | -104.9302
10.1592 | -104.9579
10.1629 | -104.9853
10.1667 | -105.0089
10.1704 | -105.0328
10.1741 | -105.0557
10.1779 | -105.0731
10.1816 | -105.0895
10.1853 | -105.1001
10.1891 | -105.1096
10.1928 | -105.1182
10.1966 | -105.1252
10.2003 | -105.1313
10.204 | -105.1338
10.2078 | -105.1383
10.2115 | -105.1417
10.2152 | -105.1372
10.219 | -105.1293
10.2227 | -105.1228
10.2264 | -105.1165
10.2302 | -105.1096
10.2339 | -105.1023
10.2376 | -105.0924
10.2414 | -105.0848
10.2451 | -105.076
10.2488 | -105.0644
10.2526 | -105.0554
10.2563 | -105.0427
10.26 | -105.0304
10.2638 | -105.0201
10.2675 | -105.0064
10.2713 | -104.9937
10.275 | -104.9837
10.2787 | -104.974
10.2825 | -104.9674
10.2862 | -104.9642
10.2899 | -104.9592
10.2937 | -104.9504
10.2974 | -104.9439
10.3011 | -104.9385
10.3049 | -104.9344
10.3086 | -104.9328
10.3123 | -104.9311
10.3161 | -104.9353
10.3198 | -104.9408
10.3235 | -104.9429
10.3273 | -104.942
10.331 | -104.9397
10.3347 | -104.9416
10.3385 | -104.942
10.3422 | -104.9402
10.346 | -104.9346
10.3497 | -104.9307
10.3534 | -104.9254
10.3572 | -104.916
10.3609 | -104.9055
10.3646 | -104.8973
10.3684 | -104.8875
10.3721 | -104.8793
10.3758 | -104.8698
10.3796 | -104.8599
10.3833 | -104.8479
10.387 | -104.8428
10.3908 | -104.8359
10.3945 | -104.8335
10.3982 | -104.8353
10.402 | -104.8379
10.4057 | -104.8407
10.4094 | -104.8461
10.4132 | -104.852
10.4169 | -104.8628
10.4207 | -104.8776
10.4244 | -104.8938
10.4281 | -104.9108
10.4319 | -104.9294
10.4356 | -104.9495
10.4393 | -104.9668
10.4431 | -104.9853
10.4468 | -105.009
10.4505 | -105.0308
10.4543 | -105.0466
10.458 | -105.0604
10.4617 | -105.0727
10.4655 | -105.0813
10.4692 | -105.0857
10.4729 | -105.0804
10.4767 | -105.0759
10.4804 | -105.0726
10.4841 | -105.0657
10.4879 | -105.0526
10.4916 | -105.0356
10.4954 | -105.0192
10.4991 | -105.0036
10.5028 | -104.9809
10.5066 | -104.9602
10.5103 | -104.9452
10.514 | -104.928
10.5178 | -104.9147
10.5215 | -104.9061
10.5252 | -104.8965
10.529 | -104.8906
10.5327 | -104.8922
10.5364 | -104.8981
10.5402 | -104.8982
10.5439 | -104.9039
10.5476 | -104.91
10.5514 | -104.915
10.5551 | -104.923
10.5588 | -104.9326
10.5626 | -104.9424
10.5663 | -104.9515
10.5701 | -104.9553
10.5738 | -104.9555
10.5775 | -104.9569
10.5813 | -104.9569
10.585 | -104.9556
10.5887 | -104.9501
10.5925 | -104.9438
10.5962 | -104.9406
10.5999 | -104.9321
10.6037 | -104.9233
10.6074 | -104.913
10.6111 | -104.9017
10.6149 | -104.8907
10.6186 | -104.8808
10.6223 | -104.8702
10.6261 | -104.8633
10.6298 | -104.8536
10.6335 | -104.8474
10.6373 | -104.8451
10.641 | -104.844
10.6448 | -104.8482
10.6485 | -104.8563
10.6522 | -104.8666
10.656 | -104.8844
10.6597 | -104.9071
10.6634 | -104.9285
10.6672 | -104.9538
10.6709 | -104.9809
10.6746 | -105.0072
10.6784 | -105.0332
10.6821 | -105.0576
10.6858 | -105.0813
10.6896 | -105.1077
10.6933 | -105.1358
10.697 | -105.1547
10.7008 | -105.1656
10.7045 | -105.1744
10.7082 | -105.1779
10.712 | -105.1771
10.7157 | -105.1723
10.7195 | -105.1628
10.7232 | -105.1505
10.7269 | -105.1373
10.7307 | -105.1194
10.7344 | -105.0985
10.7381 | -105.0756
10.7419 | -105.0536
10.7456 | -105.0318
10.7493 | -105.0089
10.7531 | -104.9905
10.7568 | -104.9733
10.7605 | -104.9569
10.7643 | -104.946
10.768 | -104.9393
10.7717 | -104.9346
10.7755 | -104.9371
10.7792 | -104.9371
10.7829 | -104.9412
10.7867 | -104.9504
10.7904 | -104.9545
10.7942 | -104.9615
10.7979 | -104.9738
10.8016 | -104.9867
10.8054 | -105.001
10.8091 | -105.0187
10.8128 | -105.0336
10.8166 | -105.0487
10.8203 | -105.0702
10.824 | -105.0893
10.8278 | -105.1046
10.8315 | -105.1259
10.8352 | -105.1467
10.839 | -105.1643
10.8427 | -105.1786
10.8464 | -105.1893
10.8502 | -105.1996
10.8539 | -105.2117
10.8576 | -105.2212
10.8614 | -105.2212
10.8651 | -105.2171
10.8689 | -105.208
10.8726 | -105.1975
10.8763 | -105.1832
10.8801 | -105.167
10.8838 | -105.1468
10.8875 | -105.1209
10.8913 | -105.0962
10.895 | -105.0762
10.8987 | -105.0535
10.9025 | -105.0336
10.9062 | -105.0147
10.9099 | -104.9937
10.9137 | -104.9754
10.9174 | -104.9634
10.9211 | -104.9509
10.9249 | -104.9404
10.9286 | -104.9321
10.9323 | -104.9277
10.9361 | -104.9253
10.9398 | -104.9215
10.9436 | -104.9198
10.9473 | -104.9163
10.951 | -104.919
10.9548 | -104.9194
10.9585 | -104.9195
10.9622 | -104.9207
10.966 | -104.9215
10.9697 | -104.918
10.9734 | -104.9186
10.9772 | -104.9189
10.9809 | -104.918
10.9846 | -104.9166
10.9884 | -104.9185
10.9921 | -104.9156
10.9958 | -104.9149
10.9996 | -104.9186
11.0033 | -104.9211
11.007 | -104.9242
11.0108 | -104.9277
11.0145 | -104.9308
11.0183 | -104.9412
11.022 | -104.957
11.0257 | -104.9687
11.0295 | -104.9806
11.0332 | -104.9972
11.0369 | -105.0147
11.0407 | -105.0321
11.0444 | -105.0523
11.0481 | -105.0773
11.0519 | -105.0994
11.0556 | -105.1168
11.0593 | -105.1383
11.0631 | -105.1594
11.0668 | -105.1818
11.0705 | -105.2062
11.0743 | -105.2315
11.078 | -105.2561
11.0817 | -105.2811
11.0855 | -105.302
11.0892 | -105.323
11.093 | -105.3432
11.0967 | -105.3602
11.1004 | -105.3763
11.1042 | -105.3902
11.1079 | -105.4017
11.1116 | -105.41
11.1154 | -105.4157
11.1191 | -105.418
11.1228 | -105.4175
11.1266 | -105.4162
11.1303 | -105.4184
11.134 | -105.4127
11.1378 | -105.399
11.1415 | -105.379
11.1452 | -105.3633
11.149 | -105.3427
11.1527 | -105.3215
11.1564 | -105.3014
11.1602 | -105.2764
11.1639 | -105.2522
11.1677 | -105.2304
11.1714 | -105.2097
11.1751 | -105.1931
11.1789 | -105.1757
11.1826 | -105.158
11.1863 | -105.1446
11.1901 | -105.1374
11.1938 | -105.1301
11.1975 | -105.1264
11.2013 | -105.1259
11.205 | -105.1312
11.2087 | -105.1401
11.2125 | -105.1535
11.2162 | -105.1704
11.2199 | -105.1879
11.2237 | -105.209
11.2274 | -105.2316
11.2311 | -105.2583
11.2349 | -105.2885
11.2386 | -105.3184
11.2424 | -105.35
11.2461 | -105.3855
11.2498 | -105.4201
11.2536 | -105.4546
11.2573 | -105.4865
11.261 | -105.5152
11.2648 | -105.5406
11.2685 | -105.564
11.2722 | -105.5822
11.276 | -105.5958
11.2797 | -105.6054
11.2834 | -105.6149
11.2872 | -105.6192
11.2909 | -105.6175
11.2946 | -105.6173
11.2984 | -105.6166
11.3021 | -105.6153
11.3058 | -105.6075
11.3096 | -105.6002
11.3133 | -105.5914
11.3171 | -105.5835
11.3208 | -105.5726
11.3245 | -105.5657
11.3283 | -105.5596
11.332 | -105.55
11.3357 | -105.5448
11.3395 | -105.5385
11.3432 | -105.5317
11.3469 | -105.5301
11.3507 | -105.5316
11.3544 | -105.53
11.3581 | -105.5278
11.3619 | -105.5252
11.3656 | -105.5231
11.3693 | -105.5183
11.3731 | -105.5136
11.3768 | -105.5098
11.3805 | -105.5042
11.3843 | -105.5028
11.388 | -105.4946
11.3918 | -105.4862
11.3955 | -105.4783
11.3992 | -105.4681
11.403 | -105.4601
11.4067 | -105.4513
11.4104 | -105.4408
11.4142 | -105.432
11.4179 | -105.4255
11.4216 | -105.422
11.4254 | -105.4185
11.4291 | -105.4155
11.4328 | -105.4133
11.4366 | -105.4145
11.4403 | -105.4176
11.444 | -105.4198
11.4478 | -105.4195
11.4515 | -105.4232
11.4552 | -105.4292
11.459 | -105.4348
11.4627 | -105.4443
11.4665 | -105.4532
11.4702 | -105.4654
11.4739 | -105.4775
11.4777 | -105.4929
11.4814 | -105.5084
11.4851 | -105.5195
11.4889 | -105.5299
11.4926 | -105.5424
11.4963 | -105.5527
11.5001 | -105.5644
11.5038 | -105.5735
11.5075 | -105.5844
11.5113 | -105.5968
11.515 | -105.6064
11.5187 | -105.6169
11.5225 | -105.6291
11.5262 | -105.6413
11.5299 | -105.6557
11.5337 | -105.6696
11.5374 | -105.6848
11.5412 | -105.6995
11.5449 | -105.7104
11.5486 | -105.7212
11.5524 | -105.7279
11.5561 | -105.7348
11.5598 | -105.739
11.5636 | -105.7382
11.5673 | -105.7365
11.571 | -105.7335
11.5748 | -105.7295
11.5785 | -105.7242
11.5822 | -105.7191
11.586 | -105.7131
11.5897 | -105.7017
11.5934 | -105.6941
11.5972 | -105.6893
11.6009 | -105.6848
11.6046 | -105.6811
11.6084 | -105.6828
11.6121 | -105.683
11.6159 | -105.6832
11.6196 | -105.6838
11.6233 | -105.6806
11.6271 | -105.6782
11.6308 | -105.6749
11.6345 | -105.6729
11.6383 | -105.6688
11.642 | -105.6626
11.6457 | -105.6517
11.6495 | -105.6416
11.6532 | -105.6326
11.6569 | -105.6228
11.6607 | -105.607
11.6644 | -105.5941
11.6681 | -105.5788
11.6719 | -105.5624
11.6756 | -105.5435
11.6793 | -105.5261
11.6831 | -105.5128
11.6868 | -105.5012
11.6906 | -105.4906
11.6943 | -105.4873
11.698 | -105.4862
11.7018 | -105.4832
11.7055 | -105.4807
11.7092 | -105.4822
11.713 | -105.4905
11.7167 | -105.4972
11.7204 | -105.5022
11.7242 | -105.511
11.7279 | -105.5165
11.7316 | -105.5248
11.7354 | -105.5316
11.7391 | -105.536
11.7428 | -105.5364
11.7466 | -105.5379
11.7503 | -105.5376
11.754 | -105.5355
11.7578 | -105.5352
11.7615 | -105.5347
11.7653 | -105.5317
11.769 | -105.53
11.7727 | -105.526
11.7765 | -105.5202
11.7802 | -105.5204
11.7839 | -105.5198
11.7877 | -105.5171
11.7914 | -105.5218
11.7951 | -105.524
11.7989 | -105.5231
11.8026 | -105.5215
11.8063 | -105.5208
11.8101 | -105.5201
11.8138 | -105.5228
11.8175 | -105.5265
11.8213 | -105.5306
11.825 | -105.5332
11.8287 | -105.5359
11.8325 | -105.5361
11.8362 | -105.538
11.84 | -105.5389
11.8437 | -105.5398
11.8474 | -105.538
11.8512 | -105.5268
11.8549 | -105.5204
11.8586 | -105.5145
11.8624 | -105.5082
11.8661 | -105.4981
11.8698 | -105.4885
11.8736 | -105.485
11.8773 | -105.4798
11.881 | -105.4713
11.8848 | -105.4599
11.8885 | -105.4548
11.8922 | -105.4493
11.896 | -105.4449
11.8997 | -105.4435
11.9034 | -105.4388
11.9072 | -105.4361
11.9109 | -105.4374
11.9147 | -105.4406
11.9184 | -105.444
11.9221 | -105.4496
11.9259 | -105.4535
11.9296 | -105.4582
11.9333 | -105.4637
11.9371 | -105.4681
11.9408 | -105.4719
11.9445 | -105.4799
11.9483 | -105.4908
11.952 | -105.4997
11.9557 | -105.5094
11.9595 | -105.5194
11.9632 | -105.5264
11.9669 | -105.5362
11.9707 | -105.5475
11.9744 | -105.5582
11.9781 | -105.5709
11.9819 | -105.5829
11.9856 | -105.5913
11.9894 | -105.5988
11.9931 | -105.606
11.9968 | -105.6131
12.0006 | -105.6209
12.0043 | -105.6292
12.008 | -105.6357
12.0118 | -105.6439
12.0155 | -105.654
12.0192 | -105.6616
12.023 | -105.6685
12.0267 | -105.675
12.0304 | -105.6788
12.0342 | -105.681
12.0379 | -105.6866
12.0416 | -105.6906
12.0454 | -105.6947
12.0491 | -105.6968
12.0528 | -105.7001
12.0566 | -105.7038
12.0603 | -105.7064
12.0641 | -105.7072
12.0678 | -105.708
12.0715 | -105.7058
12.0753 | -105.7029
12.079 | -105.6976
12.0827 | -105.693
12.0865 | -105.6867
12.0902 | -105.6799
12.0939 | -105.6705
12.0977 | -105.6598
12.1014 | -105.6519
12.1051 | -105.6425
12.1089 | -105.6289
12.1126 | -105.6209
12.1163 | -105.6136
12.1201 | -105.6034
12.1238 | -105.591
12.1275 | -105.5814
12.1313 | -105.5687
12.135 | -105.5565
12.1388 | -105.5463
12.1425 | -105.5357
12.1462 | -105.527
12.15 | -105.5178
12.1537 | -105.506
12.1574 | -105.4942
12.1612 | -105.4854
12.1649 | -105.4775
12.1686 | -105.4718
12.1724 | -105.467
12.1761 | -105.4632
12.1798 | -105.459
12.1836 | -105.4541
12.1873 | -105.4481
12.191 | -105.4458
12.1948 | -105.4509
12.1985 | -105.4572
12.2022 | -105.4636
12.206 | -105.4679
12.2097 | -105.4761
12.2135 | -105.4838
12.2172 | -105.4936
12.2209 | -105.5024
12.2247 | -105.5103
12.2284 | -105.5198
12.2321 | -105.5316
12.2359 | -105.5462
12.2396 | -105.5591
12.2433 | -105.5726
12.2471 | -105.5839
12.2508 | -105.5999
12.2545 | -105.6172
12.2583 | -105.6285
12.262 | -105.6404
12.2657 | -105.6548
12.2695 | -105.6713
12.2732 | -105.6868
12.2769 | -105.7002
12.2807 | -105.7128
12.2844 | -105.7274
12.2882 | -105.7388
12.2919 | -105.7491
12.2956 | -105.7617
12.2994 | -105.7732
12.3031 | -105.7852
12.3068 | -105.8042
12.3106 | -105.8209
12.3143 | -105.8354
12.318 | -105.8475
12.3218 | -105.8562
12.3255 | -105.8649
12.3292 | -105.8749
12.333 | -105.8869
12.3367 | -105.8976
12.3404 | -105.9092
12.3442 | -105.9154
12.3479 | -105.9174
12.3516 | -105.9194
12.3554 | -105.9209
12.3591 | -105.9214
12.3629 | -105.9209
12.3666 | -105.915
12.3703 | -105.9006
12.3741 | -105.8837
12.3778 | -105.8672
12.3815 | -105.8514
12.3853 | -105.8309
12.389 | -105.809
12.3927 | -105.7819
12.3965 | -105.7614
12.4002 | -105.7431
12.4039 | -105.7261
12.4077 | -105.7083
12.4114 | -105.6931
12.4151 | -105.6824
12.4189 | -105.6735
12.4226 | -105.6657
12.4263 | -105.6639
12.4301 | -105.6607
12.4338 | -105.661
12.4376 | -105.6669
12.4413 | -105.6746
12.445 | -105.6847
12.4488 | -105.6945
12.4525 | -105.7043
12.4562 | -105.7173
12.46 | -105.7328
12.4637 | -105.7457
12.4674 | -105.7593
12.4712 | -105.7721
12.4749 | -105.7833
12.4786 | -105.7938
12.4824 | -105.8043
12.4861 | -105.8117
12.4898 | -105.82
12.4936 | -105.8235
12.4973 | -105.8276
12.501 | -105.8322
12.5048 | -105.835
12.5085 | -105.8338
12.5123 | -105.8327
12.516 | -105.8315
12.5197 | -105.8273
12.5235 | -105.8258
12.5272 | -105.8265
12.5309 | -105.8266
12.5347 | -105.8268
12.5384 | -105.8276
12.5421 | -105.8309
12.5459 | -105.8348
12.5496 | -105.842
12.5533 | -105.8538
12.5571 | -105.8672
12.5608 | -105.8792
12.5645 | -105.8902
12.5683 | -105.9006
12.572 | -105.9161
12.5758 | -105.9307
12.5795 | -105.9427
12.5832 | -105.9534
12.587 | -105.9624
12.5907 | -105.9682
12.5944 | -105.9713
12.5982 | -105.9716
12.6019 | -105.9731
12.6056 | -105.9799
12.6094 | -105.9836
12.6131 | -105.9834
12.6168 | -105.9809
12.6206 | -105.9795
12.6243 | -105.9779
12.628 | -105.9712
12.6318 | -105.9684
12.6355 | -105.971
12.6392 | -105.9739
12.643 | -105.9762
12.6467 | -105.98
12.6505 | -105.9843
12.6542 | -105.9897
12.6579 | -105.9945
12.6617 | -106.0008
12.6654 | -106.0084
12.6691 | -106.0149
12.6729 | -106.0237
12.6766 | -106.0356
12.6803 | -106.0499
12.6841 | -106.0629
12.6878 | -106.0744
12.6915 | -106.0852
12.6953 | -106.0934
12.699 | -106.1027
12.7027 | -106.1131
12.7065 | -106.1222
12.7102 | -106.1299
12.7139 | -106.134
12.7177 | -106.1391
12.7214 | -106.1442
12.7252 | -106.1467
12.7289 | -106.1477
12.7326 | -106.1478
12.7364 | -106.1464
12.7401 | -106.1426
12.7438 | -106.141
12.7476 | -106.1384
12.7513 | -106.1359
12.755 | -106.1321
12.7588 | -106.1265
12.7625 | -106.1203
12.7662 | -106.1126
12.77 | -106.1096
12.7737 | -106.1052
12.7774 | -106.0966
12.7812 | -106.088
12.7849 | -106.0829
12.7886 | -106.0745
12.7924 | -106.0665
12.7961 | -106.0651
12.7999 | -106.0628
12.8036 | -106.0593
12.8073 | -106.0561
12.8111 | -106.0573
12.8148 | -106.0571
12.8185 | -106.0542
12.8223 | -106.0509
12.826 | -106.0508
12.8297 | -106.053
12.8335 | -106.0554
12.8372 | -106.0592
12.8409 | -106.0617
12.8447 | -106.0623
12.8484 | -106.0627
12.8521 | -106.0633
12.8559 | -106.0631
12.8596 | -106.0624
12.8633 | -106.061
12.8671 | -106.0623
12.8708 | -106.065
12.8746 | -106.0663
12.8783 | -106.0665
12.882 | -106.0643
12.8858 | -106.0625
12.8895 | -106.0627
12.8932 | -106.0654
12.897 | -106.0688
12.9007 | -106.069
12.9044 | -106.0691
12.9082 | -106.0685
12.9119 | -106.0707
12.9156 | -106.0728
12.9194 | -106.0713
12.9231 | -106.072
12.9268 | -106.0726
12.9306 | -106.0737
12.9343 | -106.0767
12.938 | -106.0803
12.9418 | -106.0802
12.9455 | -106.079
12.9493 | -106.0792
12.953 | -106.0843
12.9567 | -106.0849
12.9605 | -106.0862
12.9642 | -106.0852
12.9679 | -106.0813
12.9717 | -106.0764
12.9754 | -106.0741
12.9791 | -106.0727
12.9829 | -106.065
12.9866 | -106.0589
12.9903 | -106.0507
12.9941 | -106.0437
12.9978 | -106.0389
13.0015 | -106.032
13.0053 | -106.0278
13.009 | -106.0248
13.0127 | -106.0194
13.0165 | -106.0152
13.0202 | -106.0112
13.024 | -106.0072
13.0277 | -106.0082
13.0314 | -106.0126
13.0352 | -106.0142
13.0389 | -106.019
13.0426 | -106.0244
13.0464 | -106.0299
13.0501 | -106.0353
13.0538 | -106.0423
13.0576 | -106.0482
13.0613 | -106.0569
13.065 | -106.0661
13.0688 | -106.076
13.0725 | -106.0885
13.0762 | -106.101
13.08 | -106.1124
13.0837 | -106.1224
13.0874 | -106.1321
13.0912 | -106.1411
13.0949 | -106.1514
13.0987 | -106.1607
13.1024 | -106.1725
13.1061 | -106.183
13.1099 | -106.1882
13.1136 | -106.1907
13.1173 | -106.1938
13.1211 | -106.1956
13.1248 | -106.1924
13.1285 | -106.189
13.1323 | -106.1837
13.136 | -106.1772
13.1397 | -106.1696
13.1435 | -106.1563
13.1472 | -106.1433
13.1509 | -106.1313
13.1547 | -106.1196
13.1584 | -106.1052
13.1621 | -106.086
13.1659 | -106.0645
13.1696 | -106.0419
13.1734 | -106.022
13.1771 | -105.9998
13.1808 | -105.9774
13.1846 | -105.9536
13.1883 | -105.9275
13.192 | -105.903
13.1958 | -105.8782
13.1995 | -105.8549
13.2032 | -105.8311
13.207 | -105.8074
13.2107 | -105.7857
13.2144 | -105.7672
13.2182 | -105.7495
13.2219 | -105.7305
13.2256 | -105.7148
13.2294 | -105.7021
13.2331 | -105.6947
13.2368 | -105.6915
13.2406 | -105.6904
13.2443 | -105.6916
13.2481 | -105.6967
13.2518 | -105.705
13.2555 | -105.7171
13.2593 | -105.7327
13.263 | -105.7515
13.2667 | -105.7717
13.2705 | -105.7943
13.2742 | -105.8217
13.2779 | -105.8507
13.2817 | -105.8814
13.2854 | -105.9104
13.2891 | -105.9428
13.2929 | -105.9748
13.2966 | -106.0098
13.3003 | -106.0476
13.3041 | -106.0857
13.3078 | -106.1201
13.3115 | -106.151
13.3153 | -106.1793
13.319 | -106.2061
13.3228 | -106.2345
13.3265 | -106.2589
13.3302 | -106.2775
13.334 | -106.2923
13.3377 | -106.306
13.3414 | -106.3141
13.3452 | -106.3116
13.3489 | -106.3096
13.3526 | -106.3059
13.3564 | -106.2991
13.3601 | -106.2866
13.3638 | -106.271
13.3676 | -106.255
13.3713 | -106.2315
13.375 | -106.2035
13.3788 | -106.176
13.3825 | -106.1481
13.3862 | -106.1174
13.39 | -106.0864
13.3937 | -106.057
13.3975 | -106.0266
13.4012 | -105.9936
13.4049 | -105.96
13.4087 | -105.9271
13.4124 | -105.8972
13.4161 | -105.8663
13.4199 | -105.8341
13.4236 | -105.8032
13.4273 | -105.7768
13.4311 | -105.7515
13.4348 | -105.73
13.4385 | -105.7108
13.4423 | -105.6916
13.446 | -105.6793
13.4497 | -105.6681
13.4535 | -105.6577
13.4572 | -105.6517
13.4609 | -105.6476
13.4647 | -105.6467
13.4684 | -105.6593
13.4722 | -105.6615
13.4759 | -105.661
13.4796 | -105.6677
13.4834 | -105.6738
13.4871 | -105.6796
13.4908 | -105.6844
13.4946 | -105.6916
13.4983 | -105.6985
13.502 | -105.7043
13.5058 | -105.7111
13.5095 | -105.7199
13.5132 | -105.7244
13.517 | -105.7268
13.5207 | -105.7307
13.5244 | -105.7337
13.5282 | -105.7364
13.5319 | -105.7365
13.5356 | -105.7397
13.5394 | -105.7477
13.5431 | -105.7484
13.5469 | -105.7494
13.5506 | -105.7517
13.5543 | -105.7562
13.5581 | -105.76
13.5618 | -105.7635
13.5655 | -105.772
13.5693 | -105.7816
13.573 | -105.7908
13.5767 | -105.8019
13.5805 | -105.8164
13.5842 | -105.8299
13.5879 | -105.8459
13.5917 | -105.863
13.5954 | -105.8847
13.5991 | -105.9067
13.6029 | -105.9281
13.6066 | -105.9504
13.6103 | -105.9734
13.6141 | -105.9963
13.6178 | -106.0187
13.6216 | -106.0367
13.6253 | -106.0573
13.629 | -106.0759
13.6328 | -106.0929
13.6365 | -106.1107
13.6402 | -106.1223
13.644 | -106.1297
13.6477 | -106.1382
13.6514 | -106.1474
13.6552 | -106.1537
13.6589 | -106.1568
13.6626 | -106.158
13.6664 | -106.1575
13.6701 | -106.1606
13.6738 | -106.1633
13.6776 | -106.1555
13.6813 | -106.1419
13.685 | -106.1284
13.6888 | -106.1143
13.6925 | -106.1003
13.6963 | -106.0849
13.7 | -106.0721
13.7037 | -106.0591
13.7075 | -106.0437
13.7112 | -106.0285
13.7149 | -106.0091
13.7187 | -105.9891
13.7224 | -105.9725
13.7261 | -105.9575
13.7299 | -105.9448
13.7336 | -105.9296
13.7373 | -105.9127
13.7411 | -105.8997
13.7448 | -105.8874
13.7485 | -105.8724
13.7523 | -105.8571
13.756 | -105.8455
13.7597 | -105.8376
13.7635 | -105.8305
13.7672 | -105.8236
13.771 | -105.818
13.7747 | -105.8145
13.7784 | -105.8097
13.7822 | -105.8044
13.7859 | -105.8
13.7896 | -105.7959
13.7934 | -105.791
13.7971 | -105.7851
13.8008 | -105.7825
13.8046 | -105.7788
13.8083 | -105.7741
13.812 | -105.7688
13.8158 | -105.7613
13.8195 | -105.7577
13.8232 | -105.7497
13.827 | -105.7427
13.8307 | -105.7383
13.8344 | -105.7335
13.8382 | -105.7293
13.8419 | -105.726
13.8457 | -105.7217
13.8494 | -105.7169
13.8531 | -105.7146
13.8569 | -105.7165
13.8606 | -105.7172
13.8643 | -105.7215
13.8681 | -105.7296
13.8718 | -105.7409
13.8755 | -105.754
13.8793 | -105.7696
13.883 | -105.7857
13.8867 | -105.8063
13.8905 | -105.823
13.8942 | -105.8383
13.8979 | -105.853
13.9017 | -105.8693
13.9054 | -105.886
13.9091 | -105.9057
13.9129 | -105.9273
13.9166 | -105.9454
13.9204 | -105.9596
13.9241 | -105.9712
13.9278 | -105.9833
13.9316 | -105.995
13.9353 | -106.0057
13.939 | -106.0127
13.9428 | -106.0166
13.9465 | -106.0205
13.9502 | -106.0241
13.954 | -106.0229
13.9577 | -106.0145
13.9614 | -106.003
13.9652 | -105.9905
13.9689 | -105.9759
13.9726 | -105.9614
13.9764 | -105.9489
13.9801 | -105.9341
13.9838 | -105.9166
13.9876 | -105.8976
13.9913 | -105.8766
13.9951 | -105.8571
13.9988 | -105.8335
14.0025 | -105.8076
14.0063 | -105.7861
14.01 | -105.7656
14.0137 | -105.7468
14.0175 | -105.7292
14.0212 | -105.7097
14.0249 | -105.6922
14.0287 | -105.6755
14.0324 | -105.6625
14.0361 | -105.6533
14.0399 | -105.6459
14.0436 | -105.6441
14.0473 | -105.6445
14.0511 | -105.647
14.0548 | -105.6527
14.0585 | -105.6595
14.0623 | -105.6723
14.066 | -105.6841
14.0698 | -105.697
14.0735 | -105.7167
14.0772 | -105.7416
14.081 | -105.7696
14.0847 | -105.7992
14.0884 | -105.8281
14.0922 | -105.8575
14.0959 | -105.8883
14.0996 | -105.919
14.1034 | -105.9471
14.1071 | -105.9761
14.1108 | -106.0021
14.1146 | -106.0273
14.1183 | -106.0545
14.122 | -106.0777
14.1258 | -106.1002
14.1295 | -106.1186
14.1332 | -106.1372
14.137 | -106.1565
14.1407 | -106.1716
14.1445 | -106.1819
14.1482 | -106.1916
14.1519 | -106.1991
14.1557 | -106.2032
14.1594 | -106.2069
14.1631 | -106.2119
14.1669 | -106.217
14.1706 | -106.2189
14.1743 | -106.2179
14.1781 | -106.2177
14.1818 | -106.2192
14.1855 | -106.219
14.1893 | -106.22
14.193 | -106.2234
14.1967 | -106.2262
14.2005 | -106.2288
14.2042 | -106.2348
14.2079 | -106.2388
14.2117 | -106.2403
14.2154 | -106.2443
14.2192 | -106.2489
14.2229 | -106.2547
14.2266 | -106.2636
14.2304 | -106.2744
14.2341 | -106.2811
14.2378 | -106.2826
14.2416 | -106.2838
14.2453 | -106.2866
14.249 | -106.2908
14.2528 | -106.2941
14.2565 | -106.2967
14.2602 | -106.2971
14.264 | -106.2942
14.2677 | -106.2894
14.2714 | -106.2858
14.2752 | -106.2813
14.2789 | -106.2717
14.2826 | -106.2625
14.2864 | -106.2537
14.2901 | -106.2446
14.2939 | -106.2352
14.2976 | -106.2264
14.3013 | -106.2206
14.3051 | -106.2146
14.3088 | -106.2103
14.3125 | -106.2026
14.3163 | -106.1974
14.32 | -106.1933
14.3237 | -106.1869
14.3275 | -106.1837
14.3312 | -106.182
14.3349 | -106.1815
14.3387 | -106.1796
14.3424 | -106.1804
14.3461 | -106.186
14.3499 | -106.1911
14.3536 | -106.197
14.3573 | -106.2043
14.3611 | -106.2121
14.3648 | -106.2225
14.3686 | -106.2339
14.3723 | -106.2426
14.376 | -106.2525
14.3798 | -106.2609
14.3835 | -106.2717
14.3872 | -106.2849
14.391 | -106.298
14.3947 | -106.3102
14.3984 | -106.3193
14.4022 | -106.3255
14.4059 | -106.3341
14.4096 | -106.3416
14.4134 | -106.3467
14.4171 | -106.3503
14.4208 | -106.3498
14.4246 | -106.3469
14.4283 | -106.346
14.432 | -106.3439
14.4358 | -106.3387
14.4395 | -106.3359
14.4433 | -106.3338
14.447 | -106.3258
14.4507 | -106.3196
14.4545 | -106.3115
14.4582 | -106.3036
14.4619 | -106.2956
14.4657 | -106.2893
14.4694 | -106.2832
14.4731 | -106.2755
14.4769 | -106.265
14.4806 | -106.2516
14.4843 | -106.2415
14.4881 | -106.2328
14.4918 | -106.2235
14.4955 | -106.2132
14.4993 | -106.2032
14.503 | -106.1962
14.5067 | -106.1877
14.5105 | -106.1786
14.5142 | -106.1704
14.518 | -106.1643
14.5217 | -106.1588
14.5254 | -106.1534
14.5292 | -106.1507
14.5329 | -106.1495
14.5366 | -106.1486
14.5404 | -106.1515
14.5441 | -106.1583
14.5478 | -106.1715
14.5516 | -106.1771
14.5553 | -106.1837
14.559 | -106.1915
14.5628 | -106.1998
14.5665 | -106.2089
14.5702 | -106.2207
14.574 | -106.2318
14.5777 | -106.2423
14.5814 | -106.256
14.5852 | -106.2709
14.5889 | -106.2871
14.5927 | -106.299
14.5964 | -106.3107
14.6001 | -106.3226
14.6039 | -106.3333
14.6076 | -106.3463
14.6113 | -106.3613
14.6151 | -106.3707
14.6188 | -106.3788
14.6225 | -106.3886
14.6263 | -106.397
14.63 | -106.4021
14.6337 | -106.4058
14.6375 | -106.4045
14.6412 | -106.4009
14.6449 | -106.3996
14.6487 | -106.3958
14.6524 | -106.3899
14.6561 | -106.3863
14.6599 | -106.3794
14.6636 | -106.3725
14.6674 | -106.3636
14.6711 | -106.3539
14.6748 | -106.3456
14.6786 | -106.3386
14.6823 | -106.3326
14.686 | -106.3273
14.6898 | -106.3239
14.6935 | -106.3208
14.6972 | -106.3166
14.701 | -106.3141
14.7047 | -106.3143
14.7084 | -106.3148
14.7122 | -106.3153
14.7159 | -106.3165
14.7196 | -106.3158
14.7234 | -106.3166
14.7271 | -106.3192
14.7308 | -106.3238
14.7346 | -106.3294
14.7383 | -106.3337
14.7421 | -106.3386
14.7458 | -106.3452
14.7495 | -106.3487
14.7533 | -106.3522
14.757 | -106.3566
14.7607 | -106.3633
14.7645 | -106.3701
14.7682 | -106.3776
14.7719 | -106.3851
14.7757 | -106.3911
14.7794 | -106.3966
14.7831 | -106.3994
14.7869 | -106.3993
14.7906 | -106.4003
14.7943 | -106.3988
14.7981 | -106.3958
14.8018 | -106.3933
14.8055 | -106.3919
14.8093 | -106.387
14.813 | -106.3779
14.8168 | -106.372
14.8205 | -106.3606
14.8242 | -106.3474
14.828 | -106.3353
14.8317 | -106.3214
14.8354 | -106.3033
14.8392 | -106.2823
14.8429 | -106.2668
14.8466 | -106.2497
14.8504 | -106.2334
14.8541 | -106.2182
14.8578 | -106.2012
14.8616 | -106.1849
14.8653 | -106.171
14.869 | -106.1587
14.8728 | -106.1464
14.8765 | -106.1353
14.8802 | -106.1273
14.884 | -106.1233
14.8877 | -106.1197
14.8915 | -106.1167
14.8952 | -106.1195
14.8989 | -106.1231
14.9027 | -106.1265
14.9064 | -106.1314
14.9101 | -106.1388
14.9139 | -106.15
14.9176 | -106.1622
14.9213 | -106.1745
14.9251 | -106.1873
14.9288 | -106.201
14.9325 | -106.2186
14.9363 | -106.2332
14.94 | -106.2462
14.9437 | -106.2582
14.9475 | -106.2702
14.9512 | -106.2827
14.9549 | -106.2921
14.9587 | -106.2983
14.9624 | -106.3071
14.9662 | -106.3166
14.9699 | -106.3253
14.9736 | -106.3282
14.9774 | -106.3326
14.9811 | -106.333
14.9848 | -106.3311
14.9886 | -106.3305
14.9923 | -106.3292
14.996 | -106.3262
14.9998 | -106.3248
15.0035 | -106.3199
15.0072 | -106.314
15.011 | -106.3075
15.0147 | -106.297
15.0184 | -106.2861
15.0222 | -106.2768
15.0259 | -106.2715
15.0296 | -106.2607
15.0334 | -106.2545
15.0371 | -106.2473
15.0409 | -106.2376
15.0446 | -106.2291
15.0483 | -106.2238
15.0521 | -106.2204
15.0558 | -106.2144
15.0595 | -106.2092
15.0633 | -106.2044
15.067 | -106.1971
15.0707 | -106.1927
15.0745 | -106.1905
15.0782 | -106.1887
15.0819 | -106.1881
15.0857 | -106.1969
15.0894 | -106.2015
15.0931 | -106.2066
15.0969 | -106.2141
15.1006 | -106.2209
15.1043 | -106.2295
15.1081 | -106.2378
15.1118 | -106.2432
15.1156 | -106.2527
15.1193 | -106.2667
15.123 | -106.2785
15.1268 | -106.2901
15.1305 | -106.3019
15.1342 | -106.3141
15.138 | -106.3269
15.1417 | -106.3399
15.1454 | -106.3535
15.1492 | -106.3653
15.1529 | -106.3771
15.1566 | -106.3861
15.1604 | -106.3928
15.1641 | -106.3994
15.1678 | -106.4033
15.1716 | -106.4066
15.1753 | -106.4099
15.179 | -106.4085
15.1828 | -106.4027
15.1865 | -106.3942
15.1903 | -106.3856
15.194 | -106.3765
15.1977 | -106.3665
15.2015 | -106.3562
15.2052 | -106.3418
15.2089 | -106.3266
15.2127 | -106.3134
15.2164 | -106.297
15.2201 | -106.2785
15.2239 | -106.2637
15.2276 | -106.249
15.2313 | -106.2344
15.2351 | -106.2203
15.2388 | -106.207
15.2425 | -106.1943
15.2463 | -106.1833
15.25 | -106.1715
15.2537 | -106.1637
15.2575 | -106.1596
15.2612 | -106.1551
15.265 | -106.1522
15.2687 | -106.1512
15.2724 | -106.1514
15.2762 | -106.1528
15.2799 | -106.1575
15.2836 | -106.1676
15.2874 | -106.1772
15.2911 | -106.1872
15.2948 | -106.1976
15.2986 | -106.21
15.3023 | -106.2255
15.306 | -106.2419
15.3098 | -106.2574
15.3135 | -106.2768
15.3172 | -106.2939
15.321 | -106.3151
15.3247 | -106.3363
15.3284 | -106.3598
15.3322 | -106.383
15.3359 | -106.4032
15.3397 | -106.4201
15.3434 | -106.4372
15.3471 | -106.4531
15.3509 | -106.469
15.3546 | -106.482
15.3583 | -106.4956
15.3621 | -106.5056
15.3658 | -106.5143
15.3695 | -106.5211
15.3733 | -106.5263
15.377 | -106.5301
15.3807 | -106.5335
15.3845 | -106.5356
15.3882 | -106.5366
15.3919 | -106.5349
15.3957 | -106.5301
15.3994 | -106.525
15.4031 | -106.5217
15.4069 | -106.5198
15.4106 | -106.517
15.4144 | -106.5096
15.4181 | -106.503
15.4218 | -106.4983
15.4256 | -106.4932
15.4293 | -106.4877
15.433 | -106.4842
15.4368 | -106.4826
15.4405 | -106.4825
15.4442 | -106.4791
15.448 | -106.4767
15.4517 | -106.4763
15.4554 | -106.4757
15.4592 | -106.4755
15.4629 | -106.4746
15.4666 | -106.4754
15.4704 | -106.476
15.4741 | -106.4754
15.4778 | -106.4762
15.4816 | -106.4792
15.4853 | -106.4801
15.4891 | -106.4795
15.4928 | -106.481
15.4965 | -106.4826
15.5003 | -106.4789
15.504 | -106.4752
15.5077 | -106.4741
15.5115 | -106.4701
15.5152 | -106.4692
15.5189 | -106.4649
15.5227 | -106.4574
15.5264 | -106.4492
15.5301 | -106.4421
15.5339 | -106.4339
15.5376 | -106.425
15.5413 | -106.4147
15.5451 | -106.4073
15.5488 | -106.4009
15.5525 | -106.3939
15.5563 | -106.3865
15.56 | -106.3774
15.5638 | -106.368
15.5675 | -106.3571
15.5712 | -106.3486
15.575 | -106.3435
15.5787 | -106.3356
15.5824 | -106.3282
15.5862 | -106.3198
15.5899 | -106.3145
15.5936 | -106.3096
15.5974 | -106.3079
15.6011 | -106.3045
15.6048 | -106.3004
15.6086 | -106.2977
15.6123 | -106.2972
15.616 | -106.2986
15.6198 | -106.2993
15.6235 | -106.2916
15.6272 | -106.2891
15.631 | -106.2867
15.6347 | -106.2837
15.6385 | -106.2789
15.6422 | -106.275
15.6459 | -106.2682
15.6497 | -106.2607
15.6534 | -106.2521
15.6571 | -106.2436
15.6609 | -106.2331
15.6646 | -106.2222
15.6683 | -106.214
15.6721 | -106.2052
15.6758 | -106.1931
15.6795 | -106.1811
15.6833 | -106.1677
15.687 | -106.1542
15.6907 | -106.1396
15.6945 | -106.1245
15.6982 | -106.1099
15.7019 | -106.0947
15.7057 | -106.0797
15.7094 | -106.0653
15.7132 | -106.049
15.7169 | -106.0323
15.7206 | -106.0199
15.7244 | -106.0086
15.7281 | -105.9928
15.7318 | -105.9784
15.7356 | -105.9683
15.7393 | -105.9548
15.743 | -105.9426
15.7468 | -105.9317
15.7505 | -105.9205
15.7542 | -105.9107
15.758 | -105.9017
15.7617 | -105.8955
15.7654 | -105.8888
15.7692 | -105.8819
15.7729 | -105.8759
15.7766 | -105.8716
15.7804 | -105.8721
15.7841 | -105.8735
15.7879 | -105.8766
15.7916 | -105.8829
15.7953 | -105.8848
15.7991 | -105.8866
15.8028 | -105.8932
15.8065 | -105.9017
15.8103 | -105.9097
15.814 | -105.9172
15.8177 | -105.9272
15.8215 | -105.9355
15.8252 | -105.9439
15.8289 | -105.9529
15.8327 | -105.9628
15.8364 | -105.9714
15.8401 | -105.9796
15.8439 | -105.9877
15.8476 | -105.9961
15.8513 | -106.0058
15.8551 | -106.0113
15.8588 | -106.0126
15.8626 | -106.016
15.8663 | -106.0218
15.87 | -106.0253
15.8738 | -106.0288
15.8775 | -106.0317
15.8812 | -106.0335
15.885 | -106.0339
15.8887 | -106.0371
15.8924 | -106.041
15.8962 | -106.0424
15.8999 | -106.0452
15.9036 | -106.0507
15.9074 | -106.0568
15.9111 | -106.0647
15.9148 | -106.0744
15.9186 | -106.0832
15.9223 | -106.0929
15.926 | -106.1036
15.9298 | -106.1151
15.9335 | -106.1254
15.9373 | -106.1379
15.941 | -106.1539
15.9447 | -106.1703
15.9485 | -106.1835
15.9522 | -106.2005
15.9559 | -106.224
15.9597 | -106.2423
15.9634 | -106.2579
15.9671 | -106.2787
15.9709 | -106.3002
15.9746 | -106.3197
15.9783 | -106.3386
15.9821 | -106.3551
15.9858 | -106.37
15.9895 | -106.3857
15.9933 | -106.3979
15.997 | -106.4098
16.0007 | -106.4208
16.0045 | -106.4314
16.0082 | -106.4398
16.012 | -106.4468
16.0157 | -106.4525
16.0194 | -106.4558
16.0232 | -106.4561
16.0269 | -106.4558
16.0306 | -106.4542
16.0344 | -106.4513
16.0381 | -106.4471
16.0418 | -106.4399
16.0456 | -106.4308
16.0493 | -106.422
16.053 | -106.4124
16.0568 | -106.4025
16.0605 | -106.3903
16.0642 | -106.378
16.068 | -106.3646
16.0717 | -106.3516
16.0754 | -106.3388
16.0792 | -106.3238
16.0829 | -106.3092
16.0867 | -106.3007
16.0904 | -106.29
16.0941 | -106.2795
16.0979 | -106.2707
16.1016 | -106.2628
16.1053 | -106.256
16.1091 | -106.2517
16.1128 | -106.2474
16.1165 | -106.2421
16.1203 | -106.2371
16.124 | -106.2354
16.1277 | -106.2358
16.1315 | -106.2387
16.1352 | -106.2409
16.1389 | -106.2432
16.1427 | -106.2444
16.1464 | -106.25
16.1501 | -106.2554
16.1539 | -106.2612
16.1576 | -106.2655
16.1614 | -106.273
16.1651 | -106.2799
16.1688 | -106.2844
16.1726 | -106.2921
16.1763 | -106.3001
16.18 | -106.3086
16.1838 | -106.3159
16.1875 | -106.3207
16.1912 | -106.3279
16.195 | -106.3327
16.1987 | -106.3366
16.2024 | -106.3398
16.2062 | -106.3452
16.2099 | -106.3542
16.2136 | -106.3618
16.2174 | -106.3676
16.2211 | -106.3719
16.2248 | -106.3807
16.2286 | -106.3884
16.2323 | -106.3941
16.2361 | -106.3981
16.2398 | -106.4028
16.2435 | -106.409
16.2473 | -106.4135
16.251 | -106.4182
16.2547 | -106.4268
16.2585 | -106.4328
16.2622 | -106.4361
16.2659 | -106.4393
16.2697 | -106.4452
16.2734 | -106.4519
16.2771 | -106.4577
16.2809 | -106.4624
16.2846 | -106.4678
16.2883 | -106.474
16.2921 | -106.4802
16.2958 | -106.4859
16.2995 | -106.4895
16.3033 | -106.4939
16.307 | -106.4992
16.3108 | -106.5054
16.3145 | -106.5109
16.3182 | -106.5129
16.322 | -106.5161
16.3257 | -106.5175
16.3294 | -106.5201
16.3332 | -106.5223
16.3369 | -106.5234
16.3406 | -106.5257
16.3444 | -106.529
16.3481 | -106.5347
16.3518 | -106.5403
16.3556 | -106.5426
16.3593 | -106.5432
16.363 | -106.5454
16.3668 | -106.5494
16.3705 | -106.5534
16.3742 | -106.5576
16.378 | -106.5609
16.3817 | -106.5676
16.3855 | -106.5744
16.3892 | -106.5808
16.3929 | -106.5853
16.3967 | -106.5899
16.4004 | -106.5961
16.4041 | -106.6056
16.4079 | -106.6109
16.4116 | -106.6135
16.4153 | -106.6174
16.4191 | -106.6212
16.4228 | -106.6277
16.4265 | -106.6342
16.4303 | -106.6414
16.434 | -106.6475
16.4377 | -106.6551
16.4415 | -106.6605
16.4452 | -106.6631
16.4489 | -106.6692
16.4527 | -106.6756
16.4564 | -106.683
16.4602 | -106.6889
16.4639 | -106.6928
16.4676 | -106.6987
16.4714 | -106.7068
16.4751 | -106.7127
16.4788 | -106.7198
16.4826 | -106.729
16.4863 | -106.7379
16.49 | -106.745
16.4938 | -106.7531
16.4975 | -106.7594
16.5012 | -106.7663
16.505 | -106.7729
16.5087 | -106.7795
16.5124 | -106.7872
16.5162 | -106.7966
16.5199 | -106.8039
16.5236 | -106.8089
16.5274 | -106.8145
16.5311 | -106.8218
16.5349 | -106.8302
16.5386 | -106.8362
16.5423 | -106.8405
16.5461 | -106.8444
16.5498 | -106.8483
16.5535 | -106.8521
16.5573 | -106.8559
16.561 | -106.8606
16.5647 | -106.863
16.5685 | -106.8666
16.5722 | -106.8704
16.5759 | -106.8706
16.5797 | -106.8692
16.5834 | -106.8679
16.5871 | -106.8671
16.5909 | -106.8662
16.5946 | -106.8672
16.5983 | -106.8664
16.6021 | -106.8657
16.6058 | -106.8645
16.6096 | -106.8604
16.6133 | -106.8555
16.617 | -106.8491
16.6208 | -106.8409
16.6245 | -106.8339
16.6282 | -106.8234
16.632 | -106.8143
16.6357 | -106.8072
16.6394 | -106.7972
16.6432 | -106.788
16.6469 | -106.7809
16.6506 | -106.7715
16.6544 | -106.7634
16.6581 | -106.752
16.6618 | -106.7403
16.6656 | -106.7296
16.6693 | -106.719
16.673 | -106.7072
16.6768 | -106.6979
16.6805 | -106.6877
16.6843 | -106.677
16.688 | -106.667
16.6917 | -106.655
16.6955 | -106.6424
16.6992 | -106.6342
16.7029 | -106.6309
16.7067 | -106.6257
16.7104 | -106.621
16.7141 | -106.6174
16.7179 | -106.6119
16.7216 | -106.6068
16.7253 | -106.605
16.7291 | -106.6044
16.7328 | -106.6032
16.7365 | -106.6035
16.7403 | -106.6058
16.744 | -106.6088
16.7477 | -106.6106
16.7515 | -106.6131
16.7552 | -106.6162
16.759 | -106.6199
16.7627 | -106.6242
16.7664 | -106.6281
16.7702 | -106.6311
16.7739 | -106.6362
16.7776 | -106.6422
16.7814 | -106.6475
16.7851 | -106.6549
16.7888 | -106.661
16.7926 | -106.6672
16.7963 | -106.6728
16.8 | -106.6812
16.8038 | -106.687
16.8075 | -106.6939
16.8112 | -106.7019
16.815 | -106.71
16.8187 | -106.718
16.8224 | -106.7259
16.8262 | -106.7333
16.8299 | -106.7405
16.8337 | -106.7459
16.8374 | -106.7535
16.8411 | -106.7601
16.8449 | -106.7668
16.8486 | -106.7768
16.8523 | -106.7855
16.8561 | -106.7921
16.8598 | -106.7988
16.8635 | -106.8059
16.8673 | -106.8172
16.871 | -106.8269
16.8747 | -106.8337
16.8785 | -106.8406
16.8822 | -106.847
16.8859 | -106.8539
16.8897 | -106.8617
16.8934 | -106.8689
16.8971 | -106.8779
16.9009 | -106.8856
16.9046 | -106.8942
16.9084 | -106.9026
16.9121 | -106.9117
16.9158 | -106.9196
16.9196 | -106.9265
16.9233 | -106.934
16.927 | -106.9421
16.9308 | -106.9509
16.9345 | -106.9626
16.9382 | -106.9715
16.942 | -106.9788
16.9457 | -106.9886
16.9494 | -106.9962
16.9532 | -107.0011
16.9569 | -107.0055
16.9606 | -107.0127
16.9644 | -107.0208
16.9681 | -107.0254
16.9718 | -107.032
16.9756 | -107.0381
16.9793 | -107.0424
16.9831 | -107.045
16.9868 | -107.0463
16.9905 | -107.0482
16.9943 | -107.0488
16.998 | -107.0491
17.0017 | -107.0493
17.0055 | -107.0487
17.0092 | -107.0455
17.0129 | -107.0415
17.0167 | -107.0352
17.0204 | -107.0289
17.0241 | -107.0228
17.0279 | -107.0122
17.0316 | -107.0027
17.0353 | -106.9947
17.0391 | -106.9871
17.0428 | -106.9775
17.0465 | -106.9722
17.0503 | -106.9642
17.054 | -106.9558
17.0578 | -106.9467
17.0615 | -106.9405
17.0652 | -106.9316
17.069 | -106.9231
17.0727 | -106.9183
17.0764 | -106.9133
17.0802 | -106.9054
17.0839 | -106.8989
17.0876 | -106.8941
17.0914 | -106.8896
17.0951 | -106.8853
17.0988 | -106.8823
17.1026 | -106.8793
17.1063 | -106.8773
17.11 | -106.8742
17.1138 | -106.8699
17.1175 | -106.8664
17.1212 | -106.8638
17.125 | -106.8609
17.1287 | -106.8579
17.1325 | -106.8583
17.1362 | -106.8575
17.1399 | -106.8543
17.1437 | -106.8555
17.1474 | -106.856
17.1511 | -106.8545
17.1549 | -106.8548
17.1586 | -106.8578
17.1623 | -106.8612
17.1661 | -106.8612
17.1698 | -106.8638
17.1735 | -106.8664
17.1773 | -106.8653
17.181 | -106.8668
17.1847 | -106.87
17.1885 | -106.8723
17.1922 | -106.8747
17.1959 | -106.8777
17.1997 | -106.8809
17.2034 | -106.8833
17.2072 | -106.8827
17.2109 | -106.884
17.2146 | -106.8844
17.2184 | -106.8851
17.2221 | -106.8868
17.2258 | -106.8886
17.2296 | -106.8914
17.2333 | -106.8926
17.237 | -106.892
17.2408 | -106.8922
17.2445 | -106.893
17.2482 | -106.8934
17.252 | -106.8944
17.2557 | -106.8957
17.2594 | -106.8948
17.2632 | -106.8956
17.2669 | -106.894
17.2706 | -106.8918
17.2744 | -106.8927
17.2781 | -106.8923
17.2819 | -106.89
17.2856 | -106.8867
17.2893 | -106.8816
17.2931 | -106.8782
17.2968 | -106.8762
17.3005 | -106.8747
17.3043 | -106.8737
17.308 | -106.8718
17.3117 | -106.8733
17.3155 | -106.8735
17.3192 | -106.8698
17.3229 | -106.8681
17.3267 | -106.8674
17.3304 | -106.8678
17.3341 | -106.8696
17.3379 | -106.8698
17.3416 | -106.8685
17.3453 | -106.8696
17.3491 | -106.8699
17.3528 | -106.8684
17.3566 | -106.8688
17.3603 | -106.8684
17.364 | -106.8682
17.3678 | -106.8699
17.3715 | -106.8728
17.3752 | -106.8735
17.379 | -106.8699
17.3827 | -106.8668
17.3864 | -106.8653
17.3902 | -106.8616
17.3939 | -106.8572
17.3976 | -106.853
17.4014 | -106.8492
17.4051 | -106.8442
17.4088 | -106.8412
17.4126 | -106.8388
17.4163 | -106.8323
17.42 | -106.8275
17.4238 | -106.8228
17.4275 | -106.8161
17.4313 | -106.8115
17.435 | -106.8081
17.4387 | -106.8042
17.4425 | -106.8022
17.4462 | -106.7978
17.4499 | -106.7942
17.4537 | -106.795
17.4574 | -106.7958
17.4611 | -106.7962
17.4649 | -106.7966
17.4686 | -106.7991
17.4723 | -106.8025
17.4761 | -106.8054
17.4798 | -106.8095
17.4835 | -106.8145
17.4873 | -106.8203
17.491 | -106.827
17.4947 | -106.8356
17.4985 | -106.8436
17.5022 | -106.8504
17.506 | -106.858
17.5097 | -106.8677
17.5134 | -106.8737
17.5172 | -106.8811
17.5209 | -106.8891
17.5246 | -106.8973
17.5284 | -106.905
17.5321 | -106.9106
17.5358 | -106.9171
17.5396 | -106.9254
17.5433 | -106.9296
17.547 | -106.9321
17.5508 | -106.9356
17.5545 | -106.9404
17.5582 | -106.9434
17.562 | -106.946
17.5657 | -106.9494
17.5694 | -106.9534
17.5732 | -106.9547
17.5769 | -106.9534
17.5807 | -106.95
17.5844 | -106.949
17.5881 | -106.9482
17.5919 | -106.9464
17.5956 | -106.9429
17.5993 | -106.9394
17.6031 | -106.9341
17.6068 | -106.9273
17.6105 | -106.922
17.6143 | -106.915
17.618 | -106.9061
17.6217 | -106.8973
17.6255 | -106.8882
17.6292 | -106.8799
17.6329 | -106.8713
17.6367 | -106.8599
17.6404 | -106.8497
17.6441 | -106.8406
17.6479 | -106.8332
17.6516 | -106.8246
17.6554 | -106.8153
17.6591 | -106.8077
17.6628 | -106.802
17.6666 | -106.7989
17.6703 | -106.794
17.674 | -106.7908
17.6778 | -106.789
17.6815 | -106.7859
17.6852 | -106.7847
17.689 | -106.7852
17.6927 | -106.7865
17.6964 | -106.7916
17.7002 | -106.7959
17.7039 | -106.7987
17.7076 | -106.802
17.7114 | -106.8075
17.7151 | -106.8159
17.7188 | -106.8227
17.7226 | -106.8292
17.7263 | -106.8374
17.7301 | -106.8459
17.7338 | -106.8564
17.7375 | -106.864
17.7413 | -106.872
17.745 | -106.8809
17.7487 | -106.8919
17.7525 | -106.9038
17.7562 | -106.916
17.7599 | -106.9261
17.7637 | -106.9387
17.7674 | -106.9523
17.7711 | -106.9643
17.7749 | -106.9733
17.7786 | -106.976
17.7823 | -106.9842
17.7861 | -106.9933
17.7898 | -107.0002
17.7935 | -107.0076
17.7973 | -107.0172
17.801 | -107.0263
17.8048 | -107.0313
17.8085 | -107.034
17.8122 | -107.0379
17.816 | -107.0418
17.8197 | -107.0442
17.8234 | -107.0484
17.8272 | -107.0546
17.8309 | -107.06
17.8346 | -107.0613
17.8384 | -107.0591
17.8421 | -107.058
17.8458 | -107.0579
17.8496 | -107.0572
17.8533 | -107.0576
17.857 | -107.0603
17.8608 | -107.0612
17.8645 | -107.0624
17.8682 | -107.0648
17.872 | -107.0679
17.8757 | -107.0687
17.8795 | -107.0712
17.8832 | -107.0728
17.8869 | -107.0743
17.8907 | -107.0774
17.8944 | -107.0775
17.8981 | -107.0759
17.9019 | -107.074
17.9056 | -107.0709
17.9093 | -107.0696
17.9131 | -107.0696
17.9168 | -107.0688
17.9205 | -107.0664
17.9243 | -107.0635
17.928 | -107.0615
17.9317 | -107.0571
17.9355 | -107.0534
17.9392 | -107.05
17.9429 | -107.0501
17.9467 | -107.049
17.9504 | -107.0459
17.9542 | -107.0439
17.9579 | -107.0429
17.9616 | -107.045
17.9654 | -107.0482
17.9691 | -107.0514
17.9728 | -107.0543
17.9766 | -107.0568
17.9803 | -107.0596
17.984 | -107.0642
17.9878 | -107.0664
17.9915 | -107.0714
17.9952 | -107.0766
17.999 | -107.0807
18.0027 | -107.0837
18.0064 | -107.085
18.0102 | -107.0894
18.0139 | -107.0946
18.0176 | -107.0976
18.0214 | -107.0988
18.0251 | -107.1012
18.0289 | -107.105
18.0326 | -107.11
18.0363 | -107.1145
18.0401 | -107.1187
18.0438 | -107.1211
18.0475 | -107.1199
18.0513 | -107.1162
18.055 | -107.1131
18.0587 | -107.1097
18.0625 | -107.1081
18.0662 | -107.1061
18.0699 | -107.1025
18.0737 | -107.0974
18.0774 | -107.0899
18.0811 | -107.0829
18.0849 | -107.0737
18.0886 | -107.0665
18.0923 | -107.0632
18.0961 | -107.057
18.0998 | -107.0486
18.1036 | -107.0406
18.1073 | -107.0325
18.111 | -107.0257
18.1148 | -107.0176
18.1185 | -107.0109
18.1222 | -107.0037
18.126 | -106.9989
18.1297 | -106.9933
18.1334 | -106.9896
18.1372 | -106.9852
18.1409 | -106.9827
18.1446 | -106.9811
18.1484 | -106.9804
18.1521 | -106.9799
18.1558 | -106.9792
18.1596 | -106.9795
18.1633 | -106.9817
18.167 | -106.9832
18.1708 | -106.9842
18.1745 | -106.9868
18.1783 | -106.9911
18.182 | -106.9986
18.1857 | -107.0035
18.1895 | -107.0095
18.1932 | -107.0155
18.1969 | -107.0219
18.2007 | -107.0272
18.2044 | -107.0352
18.2081 | -107.0441
18.2119 | -107.0536
18.2156 | -107.0611
18.2193 | -107.0678
18.2231 | -107.0749
18.2268 | -107.0812
18.2305 | -107.0887
18.2343 | -107.097
18.238 | -107.1045
18.2417 | -107.1107
18.2455 | -107.1179
18.2492 | -107.1229
18.253 | -107.1246
18.2567 | -107.1269
18.2604 | -107.1278
18.2642 | -107.1269
18.2679 | -107.1242
18.2716 | -107.1212
18.2754 | -107.1181
18.2791 | -107.1152
18.2828 | -107.1101
18.2866 | -107.105
18.2903 | -107.0994
18.294 | -107.0906
18.2978 | -107.0802
18.3015 | -107.0711
18.3052 | -107.0655
18.309 | -107.0581
18.3127 | -107.0485
18.3164 | -107.0402
18.3202 | -107.0273
18.3239 | -107.0159
18.3277 | -107.0055
18.3314 | -106.9955
18.3351 | -106.9863
18.3389 | -106.9789
18.3426 | -106.9697
18.3463 | -106.9597
18.3501 | -106.9504
18.3538 | -106.9407
18.3575 | -106.9312
18.3613 | -106.9238
18.365 | -106.9178
18.3687 | -106.9101
18.3725 | -106.9001
18.3762 | -106.8925
18.3799 | -106.8859
18.3837 | -106.8811
18.3874 | -106.8775
18.3911 | -106.8758
18.3949 | -106.8732
18.3986 | -106.8729
18.4024 | -106.8742
18.4061 | -106.8762
18.4098 | -106.8792
18.4136 | -106.8821
18.4173 | -106.8883
18.421 | -106.8933
18.4248 | -106.9004
18.4285 | -106.9088
18.4322 | -106.9156
18.436 | -106.924
18.4397 | -106.9333
18.4434 | -106.9402
18.4472 | -106.95
18.4509 | -106.9596
18.4546 | -106.9687
18.4584 | -106.9789
18.4621 | -106.989
18.4658 | -106.9968
18.4696 | -107.0059
18.4733 | -107.0146
18.4771 | -107.0252
18.4808 | -107.0349
18.4845 | -107.0434
18.4883 | -107.0527
18.492 | -107.0604
18.4957 | -107.0701
18.4995 | -107.08
18.5032 | -107.0903
18.5069 | -107.0984
18.5107 | -107.105
18.5144 | -107.1137
18.5181 | -107.1201
18.5219 | -107.1279
18.5256 | -107.1337
18.5293 | -107.1358
18.5331 | -107.1397
18.5368 | -107.1416
18.5405 | -107.14
18.5443 | -107.1381
18.548 | -107.1336
18.5518 | -107.1305
18.5555 | -107.1267
18.5592 | -107.118
18.563 | -107.1111
18.5667 | -107.1033
18.5704 | -107.0946
18.5742 | -107.0828
18.5779 | -107.0695
18.5816 | -107.0571
18.5854 | -107.0436
18.5891 | -107.03
18.5928 | -107.0155
18.5966 | -107.0004
18.6003 | -106.9842
18.604 | -106.9695
18.6078 | -106.952
18.6115 | -106.9327
18.6152 | -106.915
18.619 | -106.8983
18.6227 | -106.8808
18.6265 | -106.8618
18.6302 | -106.8418
18.6339 | -106.8239
18.6377 | -106.8079
18.6414 | -106.7911
18.6451 | -106.7757
18.6489 | -106.7615
18.6526 | -106.7456
18.6563 | -106.7317
18.6601 | -106.7207
18.6638 | -106.7083
18.6675 | -106.6956
18.6713 | -106.683
18.675 | -106.6719
18.6787 | -106.6653
18.6825 | -106.6588
18.6862 | -106.6525
18.6899 | -106.6445
18.6937 | -106.6384
18.6974 | -106.6353
18.7012 | -106.6328
18.7049 | -106.6302
18.7086 | -106.63
18.7124 | -106.6286
18.7161 | -106.629
18.7198 | -106.6299
18.7236 | -106.631
18.7273 | -106.6341
18.731 | -106.6394
18.7348 | -106.6432
18.7385 | -106.6463
18.7422 | -106.6493
18.746 | -106.6525
18.7497 | -106.6547
18.7534 | -106.6585
18.7572 | -106.6636
18.7609 | -106.665
18.7646 | -106.666
18.7684 | -106.6674
18.7721 | -106.668
18.7759 | -106.6684
18.7796 | -106.669
18.7833 | -106.6707
18.7871 | -106.6696
18.7908 | -106.6671
18.7945 | -106.6644
18.7983 | -106.6601
18.802 | -106.6567
18.8057 | -106.6548
18.8095 | -106.6532
18.8132 | -106.6503
18.8169 | -106.647
18.8207 | -106.6447
18.8244 | -106.6416
18.8281 | -106.6371
18.8319 | -106.6328
18.8356 | -106.6303
18.8393 | -106.628
18.8431 | -106.6232
18.8468 | -106.6201
18.8506 | -106.6195
18.8543 | -106.6187
18.858 | -106.6229
18.8618 | -106.6225
18.8655 | -106.6232
18.8692 | -106.6253
18.873 | -106.6274
18.8767 | -106.6313
18.8804 | -106.6388
18.8842 | -106.6433
18.8879 | -106.647
18.8916 | -106.653
18.8954 | -106.659
18.8991 | -106.6664
18.9028 | -106.6729
18.9066 | -106.6796
18.9103 | -106.6885
18.914 | -106.6949
18.9178 | -106.7016
18.9215 | -106.7098
18.9253 | -106.7172
18.929 | -106.7245
18.9327 | -106.7349
18.9365 | -106.7458
18.9402 | -106.7551
18.9439 | -106.7624
18.9477 | -106.7688
18.9514 | -106.7733
18.9551 | -106.7789
18.9589 | -106.7847
18.9626 | -106.7888
18.9663 | -106.7935
18.9701 | -106.7978
18.9738 | -106.8033
18.9775 | -106.8078
18.9813 | -106.8107
18.985 | -106.8107
18.9887 | -106.8108
18.9925 | -106.8093
18.9962 | -106.8062
19 | -106.8043
19.0037 | -106.8033
19.0074 | -106.8019
19.0112 | -106.8001
19.0149 | -106.7984
19.0186 | -106.7973
19.0224 | -106.7955
19.0261 | -106.7918
19.0298 | -106.7871
19.0336 | -106.7849
19.0373 | -106.7833
19.041 | -106.7832
19.0448 | -106.7823
19.0485 | -106.7805
19.0522 | -106.7771
19.056 | -106.7764
19.0597 | -106.7807
19.0634 | -106.7823
19.0672 | -106.7815
19.0709 | -106.7831
19.0747 | -106.7851
19.0784 | -106.7856
19.0821 | -106.7882
19.0859 | -106.7914
19.0896 | -106.7977
19.0933 | -106.808
19.0971 | -106.8152
19.1008 | -106.8221
19.1045 | -106.8274
19.1083 | -106.8373
19.112 | -106.8494
19.1157 | -106.8596
19.1195 | -106.8704
19.1232 | -106.8812
19.1269 | -106.8909
19.1307 | -106.9018
19.1344 | -106.9148
19.1381 | -106.9259
19.1419 | -106.9368
19.1456 | -106.9487
19.1494 | -106.9611
19.1531 | -106.9722
19.1568 | -106.9838
19.1606 | -106.9964
19.1643 | -107.0072
19.168 | -107.0201
19.1718 | -107.0334
19.1755 | -107.0452
19.1792 | -107.0568
19.183 | -107.0668
19.1867 | -107.0769
19.1904 | -107.0872
19.1942 | -107.0963
19.1979 | -107.1061
19.2016 | -107.1151
19.2054 | -107.121
19.2091 | -107.1279
19.2128 | -107.1352
19.2166 | -107.1445
19.2203 | -107.1531
19.2241 | -107.16
19.2278 | -107.1655
19.2315 | -107.1721
19.2353 | -107.1782
19.239 | -107.1819
19.2427 | -107.1824
19.2465 | -107.1825
19.2502 | -107.1835
19.2539 | -107.1848
19.2577 | -107.1866
19.2614 | -107.187
19.2651 | -107.1878
19.2689 | -107.1875
19.2726 | -107.1853
19.2763 | -107.183
19.2801 | -107.181
19.2838 | -107.1782
19.2875 | -107.1758
19.2913 | -107.1724
19.295 | -107.1695
19.2988 | -107.1665
19.3025 | -107.1665
19.3062 | -107.1662
19.31 | -107.1637
19.3137 | -107.1601
19.3174 | -107.1581
19.3212 | -107.1564
19.3249 | -107.1568
19.3286 | -107.1571
19.3324 | -107.1559
19.3361 | -107.1579
19.3398 | -107.1599
19.3436 | -107.1593
19.3473 | -107.1581
19.351 | -107.1606
19.3548 | -107.1644
19.3585 | -107.1677
19.3622 | -107.1666
19.366 | -107.1666
19.3697 | -107.1703
19.3735 | -107.1705
19.3772 | -107.1722
19.3809 | -107.1733
19.3847 | -107.1744
19.3884 | -107.1756
19.3921 | -107.1766
19.3959 | -107.1709
19.3996 | -107.168
19.4033 | -107.1674
19.4071 | -107.1661
19.4108 | -107.1639
19.4145 | -107.1624
19.4183 | -107.1607
19.422 | -107.1571
19.4257 | -107.1533
19.4295 | -107.1518
19.4332 | -107.1483
19.4369 | -107.1424
19.4407 | -107.1368
19.4444 | -107.1301
19.4482 | -107.1239
19.4519 | -107.1177
19.4556 | -107.1098
19.4594 | -107.1033
19.4631 | -107.0979
19.4668 | -107.0902
19.4706 | -107.0847
19.4743 | -107.0792
19.478 | -107.0752
19.4818 | -107.0723
19.4855 | -107.0666
19.4892 | -107.0615
19.493 | -107.0548
19.4967 | -107.046
19.5004 | -107.0402
19.5042 | -107.0366
19.5079 | -107.0298
19.5116 | -107.0205
19.5154 | -107.0126
19.5191 | -107.006
19.5229 | -107.0008
19.5266 | -106.9959
19.5303 | -106.991
19.5341 | -106.9868
19.5378 | -106.9807
19.5415 | -106.9726
19.5453 | -106.9668
19.549 | -106.9586
19.5527 | -106.9503
19.5565 | -106.9429
19.5602 | -106.936
19.5639 | -106.9282
19.5677 | -106.9198
19.5714 | -106.9111
19.5751 | -106.9041
19.5789 | -106.8952
19.5826 | -106.8838
19.5863 | -106.8744
19.5901 | -106.8665
19.5938 | -106.86
19.5976 | -106.8537
19.6013 | -106.8459
19.605 | -106.837
19.6088 | -106.8287
19.6125 | -106.8194
19.6162 | -106.8121
19.62 | -106.8069
19.6237 | -106.8028
19.6274 | -106.797
19.6312 | -106.7895
19.6349 | -106.782
19.6386 | -106.7793
19.6424 | -106.7766
19.6461 | -106.7732
19.6498 | -106.7731
19.6536 | -106.772
19.6573 | -106.7712
19.661 | -106.7727
19.6648 | -106.7746
19.6685 | -106.7747
19.6723 | -106.7775
19.676 | -106.7837
19.6797 | -106.7897
19.6835 | -106.7953
19.6872 | -106.7988
19.6909 | -106.8006
19.6947 | -106.8037
19.6984 | -106.8075
19.7021 | -106.8129
19.7059 | -106.8172
19.7096 | -106.8171
19.7133 | -106.8195
19.7171 | -106.8245
19.7208 | -106.8299
19.7245 | -106.8333
19.7283 | -106.8358
19.732 | -106.8392
19.7357 | -106.8429
19.7395 | -106.8451
19.7432 | -106.8476
19.747 | -106.8488
19.7507 | -106.85
19.7544 | -106.8511
19.7582 | -106.8505
19.7619 | -106.8479
19.7656 | -106.8469
19.7694 | -106.8443
19.7731 | -106.8429
19.7768 | -106.8429
19.7806 | -106.8416
19.7843 | -106.8363
19.788 | -106.8325
19.7918 | -106.8306
19.7955 | -106.8275
19.7992 | -106.8244
19.803 | -106.8198
19.8067 | -106.812
19.8104 | -106.8044
19.8142 | -106.7974
19.8179 | -106.7924
19.8217 | -106.7876
19.8254 | -106.7824
19.8291 | -106.7789
19.8329 | -106.7754
19.8366 | -106.7712
19.8403 | -106.7675
19.8441 | -106.7617
19.8478 | -106.7535
19.8515 | -106.7477
19.8553 | -106.7434
19.859 | -106.7376
19.8627 | -106.7315
19.8665 | -106.7264
19.8702 | -106.7212
19.8739 | -106.7155
19.8777 | -106.7095
19.8814 | -106.7057
19.8851 | -106.702
19.8889 | -106.6955
19.8926 | -106.6908
19.8964 | -106.6866
19.9001 | -106.6831
19.9038 | -106.6798
19.9076 | -106.677
19.9113 | -106.673
19.915 | -106.6689
19.9188 | -106.663
19.9225 | -106.6603
19.9262 | -106.6594
19.93 | -106.659
19.9337 | -106.6628
19.9374 | -106.661
19.9412 | -106.6591
19.9449 | -106.6594
19.9486 | -106.6591
19.9524 | -106.6601
19.9561 | -106.6631
19.9598 | -106.6637
19.9636 | -106.6627
19.9673 | -106.6618
19.9711 | -106.6663
19.9748 | -106.671
19.9785 | -106.6749
19.9823 | -106.6787
19.986 | -106.6819
19.9897 | -106.6841
19.9935 | -106.6883
19.9972 | -106.6905
20.0009 | -106.6905
20.0047 | -106.6902
20.0084 | -106.6908
20.0121 | -106.6904
20.0159 | -106.6904
20.0196 | -106.6929
20.0233 | -106.6943
20.0271 | -106.694
20.0308 | -106.6954
20.0345 | -106.6954
20.0383 | -106.6925
20.042 | -106.6919
20.0458 | -106.6935
20.0495 | -106.6943
20.0532 | -106.6934
20.057 | -106.6901
20.0607 | -106.688
20.0644 | -106.6856
20.0682 | -106.6841
20.0719 | -106.6817
20.0756 | -106.6801
20.0794 | -106.6752
20.0831 | -106.6702
20.0868 | -106.6645
20.0906 | -106.6607
20.0943 | -106.6559
20.098 | -106.6495
20.1018 | -106.6405
20.1055 | -106.6317
20.1092 | -106.6209
20.113 | -106.6117
20.1167 | -106.6025
20.1205 | -106.5903
20.1242 | -106.5799
20.1279 | -106.5728
20.1317 | -106.5682
20.1354 | -106.5602
20.1391 | -106.5501
20.1429 | -106.5426
20.1466 | -106.5362
20.1503 | -106.528
20.1541 | -106.5182
20.1578 | -106.5095
20.1615 | -106.5047
20.1653 | -106.5012
20.169 | -106.4971
20.1727 | -106.4945
20.1765 | -106.4924
20.1802 | -106.4904
20.1839 | -106.4876
20.1877 | -106.487
20.1914 | -106.4881
20.1952 | -106.4915
20.1989 | -106.4972
20.2026 | -106.5021
20.2064 | -106.5051
20.2101 | -106.508
20.2138 | -106.5113
20.2176 | -106.5154
20.2213 | -106.5213
20.225 | -106.5315
20.2288 | -106.5404
20.2325 | -106.5476
20.2362 | -106.5578
20.24 | -106.5673
20.2437 | -106.5756
20.2474 | -106.5858
20.2512 | -106.5979
20.2549 | -106.6086
20.2586 | -106.6156
20.2624 | -106.6254
20.2661 | -106.6361
20.2699 | -106.6467
20.2736 | -106.6581
20.2773 | -106.6714
20.2811 | -106.6843
20.2848 | -106.6969
20.2885 | -106.7073
20.2923 | -106.717
20.296 | -106.7254
20.2997 | -106.7354
20.3035 | -106.7448
20.3072 | -106.7514
20.3109 | -106.7592
20.3147 | -106.7676
20.3184 | -106.7739
20.3221 | -106.7797
20.3259 | -106.7856
20.3296 | -106.7892
20.3333 | -106.7926
20.3371 | -106.7949
20.3408 | -106.7988
20.3446 | -106.8009
20.3483 | -106.799
20.352 | -106.7983
20.3558 | -106.7966
20.3595 | -106.7951
20.3632 | -106.7913
20.367 | -106.7887
20.3707 | -106.7833
20.3744 | -106.7773
20.3782 | -106.7701
20.3819 | -106.7622
20.3856 | -106.7545
20.3894 | -106.7483
20.3931 | -106.7415
20.3968 | -106.7336
20.4006 | -106.7222
20.4043 | -106.7113
20.408 | -106.7028
20.4118 | -106.6895
20.4155 | -106.6762
20.4193 | -106.6649
20.423 | -106.6525
20.4267 | -106.6397
20.4305 | -106.6277
20.4342 | -106.617
20.4379 | -106.6051
20.4417 | -106.5966
20.4454 | -106.5873
20.4491 | -106.5749
20.4529 | -106.565
20.4566 | -106.5578
20.4603 | -106.5495
20.4641 | -106.542
20.4678 | -106.5356
20.4715 | -106.5311
20.4753 | -106.5325
20.479 | -106.529
20.4827 | -106.524
20.4865 | -106.5195
20.4902 | -106.5152
20.494 | -106.515
20.4977 | -106.5143
20.5014 | -106.5139
20.5052 | -106.5148
20.5089 | -106.5139
20.5126 | -106.515
20.5164 | -106.5167
20.5201 | -106.518
20.5238 | -106.5188
20.5276 | -106.5202
20.5313 | -106.5234
20.535 | -106.5269
20.5388 | -106.5302
20.5425 | -106.5339
20.5462 | -106.5404
20.55 | -106.5451
20.5537 | -106.5494
20.5574 | -106.5553
20.5612 | -106.5603
20.5649 | -106.5628
20.5687 | -106.5668
20.5724 | -106.5697
20.5761 | -106.5738
20.5799 | -106.5788
20.5836 | -106.5845
20.5873 | -106.5914
20.5911 | -106.5941
20.5948 | -106.5953
20.5985 | -106.5979
20.6023 | -106.6004
20.606 | -106.6023
20.6097 | -106.6069
20.6135 | -106.6117
20.6172 | -106.6149
20.6209 | -106.6167
20.6247 | -106.6165
20.6284 | -106.6182
20.6321 | -106.6206
20.6359 | -106.6217
20.6396 | -106.6226
20.6434 | -106.6211
20.6471 | -106.6189
20.6508 | -106.6175
20.6546 | -106.6152
20.6583 | -106.6121
20.662 | -106.6113
20.6658 | -106.61
20.6695 | -106.6083
20.6732 | -106.6058
20.677 | -106.6024
20.6807 | -106.5976
20.6844 | -106.5931
20.6882 | -106.59
20.6919 | -106.5859
20.6956 | -106.5796
20.6994 | -106.5761
20.7031 | -106.5744
20.7068 | -106.5711
20.7106 | -106.5634
20.7143 | -106.5575
20.7181 | -106.5527
20.7218 | -106.5495
20.7255 | -106.5452
20.7293 | -106.5423
20.733 | -106.538
20.7367 | -106.5329
20.7405 | -106.5263
20.7442 | -106.5191
20.7479 | -106.5132
20.7517 | -106.51
20.7554 | -106.5047
20.7591 | -106.4964
20.7629 | -106.4879
20.7666 | -106.4801
20.7703 | -106.4736
20.7741 | -106.4661
20.7778 | -106.4578
20.7815 | -106.4506
20.7853 | -106.4442
20.789 | -106.4352
20.7928 | -106.4285
20.7965 | -106.4204
20.8002 | -106.4132
20.804 | -106.408
20.8077 | -106.4002
20.8114 | -106.391
20.8152 | -106.3829
20.8189 | -106.3745
20.8226 | -106.3645
20.8264 | -106.3563
20.8301 | -106.3496
20.8338 | -106.3411
20.8376 | -106.3335
20.8413 | -106.3257
20.845 | -106.3195
20.8488 | -106.313
20.8525 | -106.3048
20.8562 | -106.2981
20.86 | -106.2946
20.8637 | -106.2894
20.8675 | -106.2874
20.8712 | -106.2852
20.8749 | -106.2819
20.8787 | -106.2766
20.8824 | -106.2704
20.8861 | -106.2662
20.8899 | -106.2641
20.8936 | -106.2602
20.8973 | -106.2546
20.9011 | -106.2511
20.9048 | -106.2463
20.9085 | -106.2408
20.9123 | -106.2356
20.916 | -106.231
20.9197 | -106.225
20.9235 | -106.2178
20.9272 | -106.2099
20.9309 | -106.2039
20.9347 | -106.1969
20.9384 | -106.1894
20.9422 | -106.1831
20.9459 | -106.1748
20.9496 | -106.1664
20.9534 | -106.1593
20.9571 | -106.1506
20.9608 | -106.1396
20.9646 | -106.1301
20.9683 | -106.1212
20.972 | -106.1114
20.9758 | -106.1033
20.9795 | -106.0962
20.9832 | -106.0881
20.987 | -106.0803
20.9907 | -106.0744
20.9944 | -106.0671
20.9982 | -106.0584
21.0019 | -106.0499
21.0056 | -106.0441
21.0094 | -106.0379
21.0131 | -106.0365
21.0169 | -106.0324
21.0206 | -106.0272
21.0243 | -106.0228
21.0281 | -106.0196
21.0318 | -106.0172
21.0355 | -106.0147
21.0393 | -106.0152
21.043 | -106.0142
21.0467 | -106.0131
21.0505 | -106.013
21.0542 | -106.0151
21.0579 | -106.02
21.0617 | -106.0246
21.0654 | -106.0288
21.0691 | -106.0319
21.0729 | -106.0368
21.0766 | -106.0431
21.0803 | -106.0482
21.0841 | -106.0528
21.0878 | -106.0573
21.0916 | -106.0613
21.0953 | -106.0686
21.099 | -106.0754
21.1028 | -106.0816
21.1065 | -106.0868
21.1102 | -106.0916
21.114 | -106.0978
21.1177 | -106.1045
21.1214 | -106.1106
21.1252 | -106.1159
21.1289 | -106.1216
21.1326 | -106.129
21.1364 | -106.136
21.1401 | -106.1419
21.1438 | -106.1473
21.1476 | -106.1511
21.1513 | -106.153
21.155 | -106.1562
21.1588 | -106.1597
21.1625 | -106.1643
21.1663 | -106.1667
21.17 | -106.1681
21.1737 | -106.1714
21.1775 | -106.1752
21.1812 | -106.1798
21.1849 | -106.1834
21.1887 | -106.1862
21.1924 | -106.1888
21.1961 | -106.1894
21.1999 | -106.1885
21.2036 | -106.189
21.2073 | -106.1887
21.2111 | -106.1899
21.2148 | -106.1908
21.2185 | -106.192
21.2223 | -106.1925
21.226 | -106.1933
21.2297 | -106.1912
21.2335 | -106.1906
21.2372 | -106.1901
21.241 | -106.1871
21.2447 | -106.1858
21.2484 | -106.1867
21.2522 | -106.1865
21.2559 | -106.186
21.2596 | -106.1881
21.2634 | -106.1883
21.2671 | -106.1886
21.2708 | -106.1905
21.2746 | -106.1903
21.2783 | -106.1895
21.282 | -106.1898
21.2858 | -106.1912
21.2895 | -106.1912
21.2932 | -106.1924
21.297 | -106.1952
21.3007 | -106.1982
21.3044 | -106.2018
21.3082 | -106.2061
21.3119 | -106.2094
21.3157 | -106.212
21.3194 | -106.2151
21.3231 | -106.2188
21.3269 | -106.223
21.3306 | -106.2278
21.3343 | -106.2334
21.3381 | -106.2403
21.3418 | -106.2472
21.3455 | -106.2537
21.3493 | -106.2599
21.353 | -106.2646
21.3567 | -106.27
21.3605 | -106.2763
21.3642 | -106.2831
21.3679 | -106.2897
21.3717 | -106.2966
21.3754 | -106.3027
21.3791 | -106.3116
21.3829 | -106.3213
21.3866 | -106.3286
21.3904 | -106.3374
21.3941 | -106.346
21.3978 | -106.3526
21.4016 | -106.3582
21.4053 | -106.3629
21.409 | -106.37
21.4128 | -106.3789
21.4165 | -106.3872
21.4202 | -106.3946
21.424 | -106.4006
21.4277 | -106.4069
21.4314 | -106.4118
21.4352 | -106.4149
21.4389 | -106.4179
21.4426 | -106.4203
21.4464 | -106.4228
21.4501 | -106.424
21.4538 | -106.4277
21.4576 | -106.4291
21.4613 | -106.4274
21.4651 | -106.4231
21.4688 | -106.4205
21.4725 | -106.4192
21.4763 | -106.4157
21.48 | -106.4128
21.4837 | -106.4067
21.4875 | -106.4001
21.4912 | -106.3951
21.4949 | -106.3904
21.4987 | -106.387
21.5024 | -106.3838
21.5061 | -106.3786
21.5099 | -106.3738
21.5136 | -106.3681
21.5173 | -106.362
21.5211 | -106.3568
21.5248 | -106.3512
21.5285 | -106.3459
21.5323 | -106.3431
21.536 | -106.3393
21.5398 | -106.3366
21.5435 | -106.3343
21.5472 | -106.3311
21.551 | -106.3302
21.5547 | -106.3254
21.5584 | -106.3221
21.5622 | -106.3207
21.5659 | -106.3179
21.5696 | -106.3169
21.5734 | -106.3168
21.5771 | -106.3139
21.5808 | -106.3129
21.5846 | -106.3122
21.5883 | -106.3116
21.592 | -106.3139
21.5958 | -106.3167
21.5995 | -106.3165
21.6032 | -106.3172
21.607 | -106.3191
21.6107 | -106.3195
21.6145 | -106.3193
21.6182 | -106.3205
21.6219 | -106.3197
21.6257 | -106.3205
21.6294 | -106.3199
21.6331 | -106.3189
21.6369 | -106.3195
21.6406 | -106.319
21.6443 | -106.3177
21.6481 | -106.3166
21.6518 | -106.315
21.6555 | -106.3137
21.6593 | -106.3116
21.663 | -106.3087
21.6667 | -106.3062
21.6705 | -106.3045
21.6742 | -106.3033
21.6779 | -106.3021
21.6817 | -106.3005
21.6854 | -106.298
21.6892 | -106.2956
21.6929 | -106.2937
21.6966 | -106.2898
21.7004 | -106.2862
21.7041 | -106.2851
21.7078 | -106.2831
21.7116 | -106.2817
21.7153 | -106.2796
21.719 | -106.2768
21.7228 | -106.2735
21.7265 | -106.2704
21.7302 | -106.2682
21.734 | -106.2636
21.7377 | -106.2601
21.7414 | -106.2587
21.7452 | -106.2564
21.7489 | -106.2572
21.7526 | -106.257
21.7564 | -106.2579
21.7601 | -106.2585
21.7639 | -106.2577
21.7676 | -106.2563
21.7713 | -106.2554
21.7751 | -106.2529
21.7788 | -106.2519
21.7825 | -106.2517
21.7863 | -106.2515
21.79 | -106.2534
21.7937 | -106.2549
21.7975 | -106.2531
21.8012 | -106.2512
21.8049 | -106.2511
21.8087 | -106.25
21.8124 | -106.2493
21.8161 | -106.2488
21.8199 | -106.2485
21.8236 | -106.249
21.8273 | -106.2508
21.8311 | -106.2523
21.8348 | -106.2527
21.8386 | -106.2516
21.8423 | -106.2497
21.846 | -106.2491
21.8498 | -106.2499
21.8535 | -106.2492
21.8572 | -106.2499
21.861 | -106.2523
21.8647 | -106.2546
21.8684 | -106.2553
21.8722 | -106.2556
21.8759 | -106.2571
21.8796 | -106.2608
21.8834 | -106.2636
21.8871 | -106.2654
21.8908 | -106.2655
21.8946 | -106.2675
21.8983 | -106.2688
21.902 | -106.2681
21.9058 | -106.2691
21.9095 | -106.2699
21.9133 | -106.27
21.917 | -106.2714
21.9207 | -106.2747
21.9245 | -106.2775
21.9282 | -106.2797
21.9319 | -106.2828
21.9357 | -106.2858
21.9394 | -106.287
21.9431 | -106.2873
21.9469 | -106.2879
21.9506 | -106.2892
21.9543 | -106.2906
21.9581 | -106.2914
21.9618 | -106.2943
21.9655 | -106.2957
21.9693 | -106.2976
21.973 | -106.2982
21.9767 | -106.2974
21.9805 | -106.2972
21.9842 | -106.2963
21.988 | -106.2957
21.9917 | -106.297
21.9954 | -106.2958
21.9992 | -106.2933
22.0029 | -106.2932
22.0066 | -106.2957
22.0104 | -106.2963
22.0141 | -106.2954
22.0178 | -106.2964
22.0216 | -106.2977
22.0253 | -106.2997
22.029 | -106.3
22.0328 | -106.2999
22.0365 | -106.3003
22.0402 | -106.3011
22.044 | -106.3033
22.0477 | -106.3032
22.0514 | -106.3021
22.0552 | -106.3008
22.0589 | -106.2984
22.0627 | -106.2983
22.0664 | -106.2999
22.0701 | -106.2973
22.0739 | -106.2941
22.0776 | -106.2936
22.0813 | -106.294
22.0851 | -106.294
22.0888 | -106.2926
22.0925 | -106.2908
22.0963 | -106.2882
22.1 | -106.286
22.1037 | -106.2838
22.1075 | -106.2821
22.1112 | -106.2809
22.1149 | -106.2792
22.1187 | -106.277
22.1224 | -106.2745
22.1261 | -106.2726
22.1299 | -106.2711
22.1336 | -106.2703
22.1374 | -106.2683
22.1411 | -106.2645
22.1448 | -106.2609
22.1486 | -106.2581
22.1523 | -106.2563
22.156 | -106.2545
22.1598 | -106.2506
22.1635 | -106.2462
22.1672 | -106.2431
22.171 | -106.2389
22.1747 | -106.2346
22.1784 | -106.2291
22.1822 | -106.2258
22.1859 | -106.223
22.1896 | -106.2205
22.1934 | -106.2146
22.1971 | -106.2113
22.2008 | -106.2122
22.2046 | -106.2124
22.2083 | -106.2076
22.2121 | -106.2022
22.2158 | -106.1971
22.2195 | -106.1923
22.2233 | -106.189
22.227 | -106.1853
22.2307 | -106.1812
22.2345 | -106.1757
22.2382 | -106.171
22.2419 | -106.1654
22.2457 | -106.1594
22.2494 | -106.1514
22.2531 | -106.1458
22.2569 | -106.141
22.2606 | -106.1339
22.2643 | -106.1275
22.2681 | -106.1206
22.2718 | -106.113
22.2755 | -106.1036
22.2793 | -106.0942
22.283 | -106.0858
22.2868 | -106.0789
22.2905 | -106.0723
22.2942 | -106.0644
22.298 | -106.0581
22.3017 | -106.051
22.3054 | -106.0429
22.3092 | -106.0365
22.3129 | -106.0313
22.3166 | -106.0246
22.3204 | -106.0166
22.3241 | -106.01
22.3278 | -106.0033
22.3316 | -105.9956
22.3353 | -105.9888
22.339 | -105.9809
22.3428 | -105.9744
22.3465 | -105.9683
22.3502 | -105.9636
22.354 | -105.96
22.3577 | -105.9553
22.3615 | -105.9523
22.3652 | -105.95
22.3689 | -105.9482
22.3727 | -105.9461
22.3764 | -105.9437
22.3801 | -105.9414
22.3839 | -105.942
22.3876 | -105.9437
22.3913 | -105.9438
22.3951 | -105.9435
22.3988 | -105.943
22.4025 | -105.9428
22.4063 | -105.9422
22.41 | -105.9438
22.4137 | -105.9457
22.4175 | -105.9463
22.4212 | -105.9462
22.4249 | -105.9481
22.4287 | -105.9501
22.4324 | -105.9528
22.4362 | -105.9562
22.4399 | -105.96
22.4436 | -105.9629
22.4474 | -105.9665
22.4511 | -105.97
22.4548 | -105.9727
22.4586 | -105.9743
22.4623 | -105.9766
22.466 | -105.9799
22.4698 | -105.9828
22.4735 | -105.9864
22.4772 | -105.9905
22.481 | -105.9935
22.4847 | -105.9957
22.4884 | -105.998
22.4922 | -105.9989
22.4959 | -106.0018
22.4996 | -106.0051
22.5034 | -106.0071
22.5071 | -106.0095
22.5109 | -106.0116
22.5146 | -106.0117
22.5183 | -106.0118
22.5221 | -106.013
22.5258 | -106.0171
22.5295 | -106.02
22.5333 | -106.0234
22.537 | -106.0231
22.5407 | -106.0204
22.5445 | -106.0186
22.5482 | -106.0163
22.5519 | -106.0154
22.5557 | -106.0154
22.5594 | -106.016
22.5631 | -106.0156
22.5669 | -106.0155
22.5706 | -106.0169
22.5743 | -106.015
22.5781 | -106.0113
22.5818 | -106.0067
22.5856 | -106.0025
22.5893 | -105.9992
22.593 | -105.9951
22.5968 | -105.9905
22.6005 | -105.9851
22.6042 | -105.978
22.608 | -105.9709
22.6117 | -105.9665
22.6154 | -105.9615
22.6192 | -105.9561
22.6229 | -105.9517
22.6266 | -105.946
22.6304 | -105.9382
22.6341 | -105.9297
22.6378 | -105.9211
22.6416 | -105.9133
22.6453 | -105.9058
22.649 | -105.8975
22.6528 | -105.8914
22.6565 | -105.8857
22.6603 | -105.8798
22.664 | -105.8742
22.6677 | -105.8681
22.6715 | -105.8608
22.6752 | -105.8537
22.6789 | -105.8466
22.6827 | -105.8387
22.6864 | -105.8323
22.6901 | -105.8261
22.6939 | -105.8209
22.6976 | -105.8159
22.7013 | -105.8103
22.7051 | -105.8046
22.7088 | -105.7997
22.7125 | -105.7966
22.7163 | -105.794
22.72 | -105.7926
22.7237 | -105.7922
22.7275 | -105.7925
22.7312 | -105.7925
22.735 | -105.7941
22.7387 | -105.7975
22.7424 | -105.8027
22.7462 | -105.807
22.7499 | -105.811
22.7536 | -105.8152
22.7574 | -105.8185
22.7611 | -105.8241
22.7648 | -105.831
22.7686 | -105.8389
22.7723 | -105.8481
22.776 | -105.8591
22.7798 | -105.8724
22.7835 | -105.8852
22.7872 | -105.8965
22.791 | -105.9077
22.7947 | -105.9191
22.7984 | -105.9301
22.8022 | -105.9424
22.8059 | -105.9567
22.8097 | -105.9702
22.8134 | -105.985
22.8171 | -105.9985
22.8209 | -106.0113
22.8246 | -106.0237
22.8283 | -106.036
22.8321 | -106.0495
22.8358 | -106.0631
22.8395 | -106.0776
22.8433 | -106.0918
22.847 | -106.1069
22.8507 | -106.1237
22.8545 | -106.1375
22.8582 | -106.153
22.8619 | -106.1705
22.8657 | -106.1865
22.8694 | -106.2005
22.8731 | -106.2117
22.8769 | -106.2243
22.8806 | -106.2366
22.8844 | -106.2482
22.8881 | -106.2604
22.8918 | -106.2706
22.8956 | -106.2817
22.8993 | -106.2946
22.903 | -106.305
22.9068 | -106.3156
22.9105 | -106.327
22.9142 | -106.3384
22.918 | -106.349
22.9217 | -106.3568
22.9254 | -106.3635
22.9292 | -106.369
22.9329 | -106.3712
22.9366 | -106.3735
22.9404 | -106.3768
22.9441 | -106.3781
22.9478 | -106.3816
22.9516 | -106.3853
22.9553 | -106.3882
22.9591 | -106.3926
22.9628 | -106.3949
22.9665 | -106.3958
22.9703 | -106.3957
22.974 | -106.3961
22.9777 | -106.3959
22.9815 | -106.3946
22.9852 | -106.3945
22.9889 | -106.3943
22.9927 | -106.3943
22.9964 | -106.394
23.0001 | -106.3936
23.0039 | -106.3953
23.0076 | -106.3975
23.0113 | -106.3979
23.0151 | -106.399
23.0188 | -106.4019
23.0225 | -106.4049
23.0263 | -106.4073
23.03 | -106.4115
23.0338 | -106.4159
23.0375 | -106.4187
23.0412 | -106.4212
23.045 | -106.4259
23.0487 | -106.4298
23.0524 | -106.4352
23.0562 | -106.4412
23.0599 | -106.4465
23.0636 | -106.4522
23.0674 | -106.4586
23.0711 | -106.4659
23.0748 | -106.4755
23.0786 | -106.4851
23.0823 | -106.4945
23.086 | -106.5033
23.0898 | -106.5111
23.0935 | -106.5177
23.0972 | -106.5244
23.101 | -106.5319
23.1047 | -106.5384
23.1085 | -106.5457
23.1122 | -106.554
23.1159 | -106.5638
23.1197 | -106.5711
23.1234 | -106.5769
23.1271 | -106.5828
23.1309 | -106.5885
23.1346 | -106.5955
23.1383 | -106.5996
23.1421 | -106.6038
23.1458 | -106.6072
23.1495 | -106.6125
23.1533 | -106.6179
23.157 | -106.6218
23.1607 | -106.6242
23.1645 | -106.625
23.1682 | -106.6226
23.1719 | -106.623
23.1757 | -106.6199
23.1794 | -106.6179
23.1832 | -106.6163
23.1869 | -106.6146
23.1906 | -106.6134
23.1944 | -106.6118
23.1981 | -106.6109
23.2018 | -106.6104
23.2056 | -106.6066
23.2093 | -106.6033
23.213 | -106.6027
23.2168 | -106.5994
23.2205 | -106.5963
23.2242 | -106.5958
23.228 | -106.5938
23.2317 | -106.5926
23.2354 | -106.59
23.2392 | -106.5864
23.2429 | -106.5842
23.2466 | -106.582
23.2504 | -106.5807
23.2541 | -106.5799
23.2579 | -106.5769
23.2616 | -106.5765
23.2653 | -106.5757
23.2691 | -106.5742
23.2728 | -106.5751
23.2765 | -106.5753
23.2803 | -106.574
23.284 | -106.5723
23.2877 | -106.5723
23.2915 | -106.5727
23.2952 | -106.5717
23.2989 | -106.5716
23.3027 | -106.5712
23.3064 | -106.5719
23.3101 | -106.5715
23.3139 | -106.5712
23.3176 | -106.5723
23.3213 | -106.5739
23.3251 | -106.5761
23.3288 | -106.576
23.3326 | -106.5754
23.3363 | -106.5752
23.34 | -106.5751
23.3438 | -106.5759
23.3475 | -106.5763
23.3512 | -106.5772
23.355 | -106.579
23.3587 | -106.5812
23.3624 | -106.5826
23.3662 | -106.5836
23.3699 | -106.5845
23.3736 | -106.5828
23.3774 | -106.5821
23.3811 | -106.5821
23.3848 | -106.5809
23.3886 | -106.5801
23.3923 | -106.5782
23.396 | -106.5746
23.3998 | -106.5708
23.4035 | -106.5678
23.4073 | -106.5645
23.411 | -106.562
23.4147 | -106.5577
23.4185 | -106.552
23.4222 | -106.5468
23.4259 | -106.5418
23.4297 | -106.5355
23.4334 | -106.5298
23.4371 | -106.5248
23.4409 | -106.5188
23.4446 | -106.5121
23.4483 | -106.5079
23.4521 | -106.5022
23.4558 | -106.4954
23.4595 | -106.4864
23.4633 | -106.4793
23.467 | -106.4744
23.4707 | -106.4687
23.4745 | -106.463
23.4782 | -106.4564
23.482 | -106.4478
23.4857 | -106.4408
23.4894 | -106.4366
23.4932 | -106.4303
23.4969 | -106.4238
23.5006 | -106.419
23.5044 | -106.415
23.5081 | -106.4113
23.5118 | -106.406
23.5156 | -106.4024
23.5193 | -106.3983
23.523 | -106.3936
23.5268 | -106.3887
23.5305 | -106.3835
23.5342 | -106.3813
23.538 | -106.3787
23.5417 | -106.3764
23.5454 | -106.3729
23.5492 | -106.3723
23.5529 | -106.3713
23.5567 | -106.3706
23.5604 | -106.3704
23.5641 | -106.3708
23.5679 | -106.3728
23.5716 | -106.3736
23.5753 | -106.3741
23.5791 | -106.3771
23.5828 | -106.3805
23.5865 | -106.3841
23.5903 | -106.3912
23.594 | -106.3962
23.5977 | -106.4028
23.6015 | -106.4116
23.6052 | -106.4194
23.6089 | -106.4273
23.6127 | -106.4365
23.6164 | -106.448
23.6201 | -106.4584
23.6239 | -106.4673
23.6276 | -106.478
23.6314 | -106.4886
23.6351 | -106.5005
23.6388 | -106.5109
23.6426 | -106.5225
23.6463 | -106.5342
23.65 | -106.5452
23.6538 | -106.5575
23.6575 | -106.5688
23.6612 | -106.5811
23.665 | -106.5944
23.6687 | -106.6074
23.6724 | -106.6202
23.6762 | -106.6316
23.6799 | -106.6402
23.6836 | -106.65
23.6874 | -106.6587
23.6911 | -106.6659
23.6948 | -106.6734
23.6986 | -106.6813
23.7023 | -106.6877
23.7061 | -106.6879
23.7098 | -106.6948
23.7135 | -106.6992
23.7173 | -106.7016
23.721 | -106.7047
23.7247 | -106.7086
23.7285 | -106.7119
23.7322 | -106.7145
23.7359 | -106.7171
23.7397 | -106.7161
23.7434 | -106.7147
23.7471 | -106.7139
23.7509 | -106.7137
23.7546 | -106.7129
23.7583 | -106.7099
23.7621 | -106.7064
23.7658 | -106.7024
23.7695 | -106.698
23.7733 | -106.6917
23.777 | -106.6876
23.7808 | -106.6827
23.7845 | -106.676
23.7882 | -106.6691
23.792 | -106.6625
23.7957 | -106.657
23.7994 | -106.65
23.8032 | -106.6426
23.8069 | -106.6357
23.8106 | -106.6292
23.8144 | -106.6229
23.8181 | -106.6178
23.8218 | -106.6115
23.8256 | -106.6026
23.8293 | -106.5949
23.833 | -106.5871
23.8368 | -106.5803
23.8405 | -106.5739
23.8442 | -106.5682
23.848 | -106.5632
23.8517 | -106.5566
23.8555 | -106.5512
23.8592 | -106.546
23.8629 | -106.5415
23.8667 | -106.5365
23.8704 | -106.532
23.8741 | -106.5277
23.8779 | -106.5225
23.8816 | -106.5186
23.8853 | -106.5136
23.8891 | -106.5094
23.8928 | -106.5055
23.8965 | -106.5011
23.9003 | -106.4984
23.904 | -106.4952
23.9077 | -106.4922
23.9115 | -106.489
23.9152 | -106.4842
23.9189 | -106.4806
23.9227 | -106.478
23.9264 | -106.4745
23.9302 | -106.4723
23.9339 | -106.4694
23.9376 | -106.4659
23.9414 | -106.4621
23.9451 | -106.4583
23.9488 | -106.4552
23.9526 | -106.4513
23.9563 | -106.4484
23.96 | -106.448
23.9638 | -106.4463
23.9675 | -106.443
23.9712 | -106.4387
23.975 | -106.4345
23.9787 | -106.4308
23.9824 | -106.4272
23.9862 | -106.424
23.9899 | -106.4199
23.9936 | -106.4179
23.9974 | -106.4163
24.0011 | -106.4139
24.0049 | -106.4117
24.0086 | -106.4103
24.0123 | -106.4079
24.0161 | -106.406
24.0198 | -106.4035
24.0235 | -106.403
24.0273 | -106.4021
24.031 | -106.402
24.0347 | -106.402
24.0385 | -106.4024
24.0422 | -106.4027
24.0459 | -106.401
24.0497 | -106.3973
24.0534 | -106.3966
24.0571 | -106.3974
24.0609 | -106.3972
24.0646 | -106.395
24.0684 | -106.3939
24.0721 | -106.3939
24.0758 | -106.3943
24.0796 | -106.3971
24.0833 | -106.3995
24.087 | -106.3996
24.0908 | -106.3983
24.0945 | -106.3986
24.0982 | -106.3997
24.102 | -106.3985
24.1057 | -106.3974
24.1094 | -106.3968
24.1132 | -106.3971
24.1169 | -106.3954
24.1206 | -106.3934
24.1244 | -106.3936
24.1281 | -106.3952
24.1318 | -106.3968
24.1356 | -106.3971
24.1393 | -106.3955
24.1431 | -106.3938
24.1468 | -106.3938
24.1505 | -106.3938
24.1543 | -106.3942
24.158 | -106.3939
24.1617 | -106.3925
24.1655 | -106.3928
24.1692 | -106.3941
24.1729 | -106.3958
24.1767 | -106.3952
24.1804 | -106.3946
24.1841 | -106.3945
24.1879 | -106.3933
24.1916 | -106.3937
24.1953 | -106.3949
24.1991 | -106.3952
24.2028 | -106.396
24.2065 | -106.3973
24.2103 | -106.3965
24.214 | -106.3954
24.2178 | -106.3974
24.2215 | -106.3974
24.2252 | -106.3987
24.229 | -106.4001
24.2327 | -106.4011
24.2364 | -106.4027
24.2402 | -106.4032
24.2439 | -106.4052
24.2476 | -106.4062
24.2514 | -106.405
24.2551 | -106.4049
24.2588 | -106.4011
24.2626 | -106.388
24.2663 | -106.3657
24.27 | -106.328
24.2738 | -106.2932
24.2775 | -106.2584
24.2812 | -106.2299
24.285 | -106.2064
24.2887 | -106.1828
24.2925 | -106.1572
24.2962 | -106.1287
24.2999 | -106.0989
24.3037 | -106.0675
24.3074 | -106.0341
24.3111 | -105.999
24.3149 | -105.9614
24.3186 | -105.9218
24.3223 | -105.8794
24.3261 | -105.8352
24.3298 | -105.7896
24.3335 | -105.7419
24.3373 | -105.6909
24.341 | -105.6396
24.3447 | -105.5855
24.3485 | -105.5293
24.3522 | -105.4743
24.3559 | -105.42
24.3597 | -105.3643
24.3634 | -105.3095
24.3672 | -105.2533
24.3709 | -105.1992
24.3746 | -105.1459
24.3784 | -105.0929
24.3821 | -105.0406
24.3858 | -104.9888
24.3896 | -104.9409
24.3933 | -104.8965
24.397 | -104.8522
24.4008 | -104.8099
24.4045 | -104.7707
24.4082 | -104.7346
24.412 | -104.7015
24.4157 | -104.6713
24.4194 | -104.6448
24.4232 | -104.6214
24.4269 | -104.6009
24.4306 | -104.5844
24.4344 | -104.5722
24.4381 | -104.5643
24.4419 | -104.5591
24.4456 | -104.5588
24.4493 | -104.5625
24.4531 | -104.5683
24.4568 | -104.5768
24.4605 | -104.5886
24.4643 | -104.6036
24.468 | -104.6217
24.4717 | -104.6426
24.4755 | -104.6666
24.4792 | -104.6939
24.4829 | -104.7238
24.4867 | -104.7558
24.4904 | -104.7893
24.4941 | -104.826
24.4979 | -104.8638
24.5016 | -104.9037
24.5053 | -104.9452
24.5091 | -104.9867
24.5128 | -105.0293
24.5166 | -105.0726
24.5203 | -105.1156
24.524 | -105.157
24.5278 | -105.1991
24.5315 | -105.24
24.5352 | -105.2802
24.539 | -105.3191
24.5427 | -105.357
24.5464 | -105.3937
24.5502 | -105.4285
24.5539 | -105.4607
24.5576 | -105.4926
24.5614 | -105.5205
24.5651 | -105.546
24.5688 | -105.5655
24.5726 | -105.5695
24.5763 | -105.5243
24.58 | -105.4353
24.5838 | -105.3574
24.5875 | -105.293
24.5913 | -105.2422
24.595 | -105.1984
24.5987 | -105.1526
24.6025 | -105.102
24.6062 | -105.0493
24.6099 | -104.9955
24.6137 | -104.9402
24.6174 | -104.8856
24.6211 | -104.819
24.6249 | -104.7442
24.6286 | -104.6578
24.6323 | -104.5641
24.6361 | -104.4658
24.6398 | -104.3662
24.6435 | -104.2649
24.6473 | -104.1599
24.651 | -104.0542
24.6547 | -103.948
24.6585 | -103.842
24.6622 | -103.7358
24.666 | -103.6296
24.6697 | -103.5242
24.6734 | -103.4208
24.6772 | -103.3187
24.6809 | -103.2196
24.6846 | -103.1232
24.6884 | -103.0289
24.6921 | -102.9386
24.6958 | -102.8527
24.6996 | -102.7688
24.7033 | -102.6879
24.707 | -102.6131
24.7108 | -102.543
24.7145 | -102.4778
24.7182 | -102.4176
24.722 | -102.3628
24.7257 | -102.3135
24.7294 | -102.2705
24.7332 | -102.2339
24.7369 | -102.2027
24.7407 | -102.1774
24.7444 | -102.1582
24.7481 | -102.1447
24.7519 | -102.1376
24.7556 | -102.1376
24.7593 | -102.1423
24.7631 | -102.1541
24.7668 | -102.1723
24.7705 | -102.1958
24.7743 | -102.2259
24.778 | -102.2626
24.7817 | -102.3047
24.7855 | -102.3914
24.7892 | -102.4412
24.7929 | -102.4965
24.7967 | -102.5562
24.8004 | -102.6201
24.8041 | -102.6889
24.8079 | -102.7617
24.8116 | -102.8388
24.8154 | -102.9196
24.8191 | -103.0033
24.8228 | -103.0894
24.8266 | -103.177
24.8303 | -103.266
24.834 | -103.3569
24.8378 | -103.4491
24.8415 | -103.5427
24.8452 | -103.6365
24.849 | -103.7304
24.8527 | -103.8247
24.8564 | -103.9175
24.8602 | -104.0091
24.8639 | -104.0994
24.8676 | -104.1879
24.8714 | -104.2734
24.8751 | -104.3573
24.8788 | -104.4394
24.8826 | -104.5183
24.8863 | -104.5923
24.8901 | -104.6638
24.8938 | -104.7324
24.8975 | -104.7966
24.9013 | -104.8583
24.905 | -104.9174
24.9087 | -104.9711
24.9125 | -105.0202
24.9162 | -105.068
24.9199 | -105.1145
24.9237 | -105.167
24.9274 | -105.2499
24.9311 | -105.3547
24.9349 | -105.4284
24.9386 | -105.4713
24.9423 | -105.4887
24.9461 | -105.4921
24.9498 | -105.4924
24.9535 | -105.4902
24.9573 | -105.4865
24.961 | -105.484
24.9648 | -105.4796
24.9685 | -105.4759
24.9722 | -105.4711
24.976 | -105.4654
24.9797 | -105.4597
24.9834 | -105.4554
24.9872 | -105.4511
24.9909 | -105.4469
24.9946 | -105.4424
24.9984 | -105.4384
25.0021 | -105.4328
25.0058 | -105.4285
25.0096 | -105.4217
25.0133 | -105.4156
25.017 | -105.41
25.0208 | -105.4046
25.0245 | -105.3989
25.0282 | -105.3921
25.032 | -105.3832
25.0357 | -105.376
25.0395 | -105.3692
25.0432 | -105.3625
25.0469 | -105.3535
25.0507 | -105.346
25.0544 | -105.3399
25.0581 | -105.3316
25.0619 | -105.3225
25.0656 | -105.3153
25.0693 | -105.3069
25.0731 | -105.298
25.0768 | -105.2905
25.0805 | -105.2806
25.0843 | -105.2705
25.088 | -105.2609
25.0917 | -105.2511
25.0955 | -105.2431
25.0992 | -105.2373
25.1029 | -105.2298
25.1067 | -105.2207
25.1104 | -105.2132
25.1142 | -105.2037
25.1179 | -105.1938
25.1216 | -105.1847
25.1254 | -105.1746
25.1291 | -105.1648
25.1328 | -105.1556
25.1366 | -105.1455
25.1403 | -105.1337
25.144 | -105.1236
25.1478 | -105.1136
25.1515 | -105.1042
25.1552 | -105.0926
25.159 | -105.0813
25.1627 | -105.0733
25.1664 | -105.0646
25.1702 | -105.0552
25.1739 | -105.0451
25.1776 | -105.0357
25.1814 | -105.0272
25.1851 | -105.0197
25.1889 | -105.0124
25.1926 | -105.0022
25.1963 | -104.9897
25.2001 | -104.9774
25.2038 | -104.9666
25.2075 | -104.9558
25.2113 | -104.947
25.215 | -104.9373
25.2187 | -104.9265
25.2225 | -104.9168
25.2262 | -104.9078
25.2299 | -104.8991
25.2337 | -104.8895
25.2374 | -104.8805
25.2411 | -104.8724
25.2449 | -104.8632
25.2486 | -104.8542
25.2523 | -104.8456
25.2561 | -104.8368
25.2598 | -104.8264
25.2636 | -104.8178
25.2673 | -104.8102
25.271 | -104.8025
25.2748 | -104.7947
25.2785 | -104.7861
25.2822 | -104.7796
25.286 | -104.7719
25.2897 | -104.7642
25.2934 | -104.7574
25.2972 | -104.7509
25.3009 | -104.7446
25.3046 | -104.7397
25.3084 | -104.7354
25.3121 | -104.731
25.3158 | -104.7267
25.3196 | -104.7251
25.3233 | -104.7287
25.327 | -104.7269
25.3308 | -104.725
25.3345 | -104.7241
25.3383 | -104.724
25.342 | -104.7227
25.3457 | -104.7243
25.3495 | -104.7259
25.3532 | -104.7275
25.3569 | -104.7299
25.3607 | -104.7346
25.3644 | -104.7393
25.3681 | -104.7448
25.3719 | -104.7494
25.3756 | -104.7543
25.3793 | -104.7602
25.3831 | -104.7655
25.3868 | -104.7716
25.3905 | -104.7777
25.3943 | -104.7844
25.398 | -104.7924
25.4017 | -104.7977
25.4055 | -104.804
25.4092 | -104.8103
25.413 | -104.8175
25.4167 | -104.8247
25.4204 | -104.832
25.4242 | -104.839
25.4279 | -104.8477
25.4316 | -104.8557
25.4354 | -104.8639
25.4391 | -104.8739
25.4428 | -104.8832
25.4466 | -104.8908
25.4503 | -104.8987
25.454 | -104.9064
25.4578 | -104.9131
25.4615 | -104.921
25.4652 | -104.9284
25.469 | -104.9351
25.4727 | -104.9446
25.4764 | -104.9539
25.4802 | -104.9634
25.4839 | -104.9731
25.4877 | -104.9819
25.4914 | -104.9907
25.4951 | -105.0002
25.4989 | -105.0101
25.5026 | -105.0185
25.5063 | -105.0268
25.5101 | -105.0346
25.5138 | -105.0457
25.5175 | -105.0554
25.5213 | -105.0633
25.525 | -105.0713
25.5287 | -105.079
25.5325 | -105.0885
25.5362 | -105.0986
25.5399 | -105.1095
25.5437 | -105.1184
25.5474 | -105.1258
25.5511 | -105.1371
25.5549 | -105.1513
25.5586 | -105.1639
25.5624 | -105.1747
25.5661 | -105.185
25.5698 | -105.1946
25.5736 | -105.2056
25.5773 | -105.2158
25.581 | -105.2245
25.5848 | -105.2341
25.5885 | -105.2438
25.5922 | -105.2511
25.596 | -105.2594
25.5997 | -105.2689
25.6034 | -105.2777
25.6072 | -105.2866
25.6109 | -105.2955
25.6146 | -105.3031
25.6184 | -105.3133
25.6221 | -105.3231
25.6258 | -105.3316
25.6296 | -105.3378
25.6333 | -105.3436
25.6371 | -105.3504
25.6408 | -105.3571
25.6445 | -105.3661
25.6483 | -105.3731
25.652 | -105.3792
25.6557 | -105.3842
25.6595 | -105.388
25.6632 | -105.3906
25.6669 | -105.3938
25.6707 | -105.3961
25.6744 | -105.3989
25.6781 | -105.4012
25.6819 | -105.4025
25.6856 | -105.4047
25.6893 | -105.4077
25.6931 | -105.4104
25.6968 | -105.4098
25.7005 | -105.4095
25.7043 | -105.4092
25.708 | -105.4071
25.7118 | -105.4051
25.7155 | -105.403
25.7192 | -105.4021
25.723 | -105.3992
25.7267 | -105.3963
25.7304 | -105.3932
25.7342 | -105.3932
25.7379 | -105.3931
25.7416 | -105.392
25.7454 | -105.3901
25.7491 | -105.389
25.7528 | -105.3902
25.7566 | -105.3892
25.7603 | -105.3866
25.764 | -105.3875
25.7678 | -105.3879
25.7715 | -105.3859
25.7752 | -105.3832
25.779 | -105.383
25.7827 | -105.3824
25.7865 | -105.3825
25.7902 | -105.3813
25.7939 | -105.3809
25.7977 | -105.3805
25.8014 | -105.3829
25.8051 | -105.3868
25.8089 | -105.3918
25.8126 | -105.396
25.8163 | -105.4002
25.8201 | -105.4056
25.8238 | -105.4103
25.8275 | -105.4159
25.8313 | -105.4234
25.835 | -105.428
25.8387 | -105.433
25.8425 | -105.4386
25.8462 | -105.4451
25.8499 | -105.4518
25.8537 | -105.4595
25.8574 | -105.4651
25.8612 | -105.4727
25.8649 | -105.4806
25.8686 | -105.4894
25.8724 | -105.4977
25.8761 | -105.5056
25.8798 | -105.5124
25.8836 | -105.5218
25.8873 | -105.5293
25.891 | -105.5382
25.8948 | -105.5481
25.8985 | -105.5569
25.9022 | -105.5652
25.906 | -105.572
25.9097 | -105.5821
25.9134 | -105.5923
25.9172 | -105.6004
25.9209 | -105.6071
25.9246 | -105.6143
25.9284 | -105.6208
25.9321 | -105.6275
25.9359 | -105.6333
25.9396 | -105.6395
25.9433 | -105.6474
25.9471 | -105.6549
25.9508 | -105.6617
25.9545 | -105.6698
25.9583 | -105.6791
25.962 | -105.6854
25.9657 | -105.6903
25.9695 | -105.6948
25.9732 | -105.6986
25.9769 | -105.7028
25.9807 | -105.7078
25.9844 | -105.7117
25.9881 | -105.7151
25.9919 | -105.72
25.9956 | -105.7243
25.9993 | -105.7302
26.0031 | -105.7351
26.0068 | -105.7385
26.0106 | -105.74
26.0143 | -105.743
26.018 | -105.7455
26.0218 | -105.7476
26.0255 | -105.7501
26.0292 | -105.7526
26.033 | -105.7541
26.0367 | -105.7568
26.0404 | -105.7583
26.0442 | -105.7583
26.0479 | -105.7584
26.0516 | -105.7592
26.0554 | -105.7593
26.0591 | -105.7602
26.0628 | -105.7614
26.0666 | -105.7621
26.0703 | -105.7628
26.074 | -105.7636
26.0778 | -105.7643
26.0815 | -105.7657
26.0853 | -105.7666
26.089 | -105.7682
26.0927 | -105.7692
26.0965 | -105.7722
26.1002 | -105.7735
26.1039 | -105.7752
26.1077 | -105.777
26.1114 | -105.7804
26.1151 | -105.7831
26.1189 | -105.7843
26.1226 | -105.7874
26.1263 | -105.791
26.1301 | -105.7939
26.1338 | -105.796
26.1375 | -105.801
26.1413 | -105.8065
26.145 | -105.8097
26.1487 | -105.8136
26.1525 | -105.8185
26.1562 | -105.8238
26.16 | -105.8285
26.1637 | -105.8338
26.1674 | -105.8366
26.1712 | -105.8391
26.1749 | -105.8422
26.1786 | -105.8473
26.1824 | -105.8516
26.1861 | -105.8559
26.1898 | -105.8632
26.1936 | -105.8699
26.1973 | -105.876
26.201 | -105.8813
26.2048 | -105.8854
26.2085 | -105.8905
26.2122 | -105.8949
26.216 | -105.8981
26.2197 | -105.9012
26.2234 | -105.9079
26.2272 | -105.9133
26.2309 | -105.9187
26.2347 | -105.9227
26.2384 | -105.9263
26.2421 | -105.9303
26.2459 | -105.9343
26.2496 | -105.9388
26.2533 | -105.9419
26.2571 | -105.9458
26.2608 | -105.9498
26.2645 | -105.9536
26.2683 | -105.9577
26.272 | -105.9608
26.2757 | -105.9645
26.2795 | -105.9676
26.2832 | -105.9705
26.2869 | -105.974
26.2907 | -105.9772
26.2944 | -105.9804
26.2981 | -105.9852
26.3019 | -105.9888
26.3056 | -105.9918
26.3094 | -105.9954
26.3131 | -105.995
26.3168 | -105.9953
26.3206 | -105.9967
26.3243 | -105.9956
26.328 | -105.996
26.3318 | -105.9975
26.3355 | -105.999
26.3392 | -106.0006
26.343 | -106.0003
26.3467 | -105.9991
26.3504 | -105.998
26.3542 | -105.9948
26.3579 | -105.992
26.3616 | -105.9879
26.3654 | -105.9847
26.3691 | -105.9798
26.3728 | -105.9757
26.3766 | -105.9725
26.3803 | -105.969
26.3841 | -105.9679
26.3878 | -105.9659
26.3915 | -105.9631
26.3953 | -105.9593
26.399 | -105.9555
26.4027 | -105.9519
26.4065 | -105.9479
26.4102 | -105.9441
26.4139 | -105.9413
26.4177 | -105.9382
26.4214 | -105.9349
26.4251 | -105.9301
26.4289 | -105.9254
26.4326 | -105.9215
26.4363 | -105.9193
26.4401 | -105.9165
26.4438 | -105.9121
26.4475 | -105.9081
26.4513 | -105.9047
26.455 | -105.9017
26.4588 | -105.899
26.4625 | -105.8949
26.4662 | -105.8896
26.47 | -105.8864
26.4737 | -105.8822
26.4774 | -105.8774
26.4812 | -105.8737
26.4849 | -105.8702
26.4886 | -105.8688
26.4924 | -105.8684
26.4961 | -105.8679
26.4998 | -105.8682
26.5036 | -105.8696
26.5073 | -105.8729
26.511 | -105.8742
26.5148 | -105.8741
26.5185 | -105.8753
26.5222 | -105.8771
26.526 | -105.8778
26.5297 | -105.8801
26.5335 | -105.8815
26.5372 | -105.8821
26.5409 | -105.8851
26.5447 | -105.8886
26.5484 | -105.893
26.5521 | -105.8974
26.5559 | -105.9019
26.5596 | -105.907
26.5633 | -105.9102
26.5671 | -105.9146
26.5708 | -105.9203
26.5745 | -105.9265
26.5783 | -105.9336
26.582 | -105.9415
26.5857 | -105.9485
26.5895 | -105.9566
26.5932 | -105.9637
26.5969 | -105.9691
26.6007 | -105.9749
26.6044 | -105.981
26.6082 | -105.9882
26.6119 | -105.9941
26.6156 | -106.0007
26.6194 | -106.01
26.6231 | -106.0182
26.6268 | -106.0256
26.6306 | -106.033
26.6343 | -106.0407
26.638 | -106.0484
26.6418 | -106.0557
26.6455 | -106.0622
26.6492 | -106.0682
26.653 | -106.0746
26.6567 | -106.0828
26.6604 | -106.0892
26.6642 | -106.0962
26.6679 | -106.1035
26.6716 | -106.1093
26.6754 | -106.1161
26.6791 | -106.124
26.6829 | -106.1311
26.6866 | -106.1388
26.6903 | -106.1462
26.6941 | -106.1515
26.6978 | -106.1566
26.7015 | -106.1606
26.7053 | -106.1646
26.709 | -106.1677
26.7127 | -106.1717
26.7165 | -106.1775
26.7202 | -106.1832
26.7239 | -106.1883
26.7277 | -106.1937
26.7314 | -106.1949
26.7351 | -106.1974
26.7389 | -106.2011
26.7426 | -106.2045
26.7463 | -106.2057
26.7501 | -106.2051
26.7538 | -106.2051
26.7576 | -106.204
26.7613 | -106.204
26.765 | -106.2046
26.7688 | -106.2026
26.7725 | -106.1997
26.7762 | -106.1961
26.78 | -106.1935
26.7837 | -106.1906
26.7874 | -106.188
26.7912 | -106.1834
26.7949 | -106.18
26.7986 | -106.1782
26.8024 | -106.1762
26.8061 | -106.1735
26.8098 | -106.1702
26.8136 | -106.1661
26.8173 | -106.1621
26.821 | -106.1588
26.8248 | -106.1537
26.8285 | -106.1476
26.8323 | -106.1435
26.836 | -106.1401
26.8397 | -106.1379
26.8435 | -106.1351
26.8472 | -106.1299
26.8509 | -106.1277
26.8547 | -106.1251
26.8584 | -106.1211
26.8621 | -106.1179
26.8659 | -106.1141
26.8696 | -106.1091
26.8733 | -106.1037
26.8771 | -106.0991
26.8808 | -106.0938
26.8845 | -106.0901
26.8883 | -106.0873
26.892 | -106.0843
26.8957 | -106.0798
26.8995 | -106.0754
26.9032 | -106.0711
26.907 | -106.0676
26.9107 | -106.0653
26.9144 | -106.0643
26.9182 | -106.0614
26.9219 | -106.0576
26.9256 | -106.0525
26.9294 | -106.0483
26.9331 | -106.045
26.9368 | -106.0413
26.9406 | -106.0404
26.9443 | -106.0373
26.948 | -106.0348
26.9518 | -106.0314
26.9555 | -106.0282
26.9592 | -106.0258
26.963 | -106.0221
26.9667 | -106.0189
26.9704 | -106.0159
26.9742 | -106.0129
26.9779 | -106.0103
26.9817 | -106.0083
26.9854 | -106.0054
26.9891 | -106.0011
26.9929 | -105.9979
26.9966 | -105.9958
27.0003 | -105.9918
27.0041 | -105.9889
27.0078 | -105.9867
27.0115 | -105.9847
27.0153 | -105.9817
27.019 | -105.9801
27.0227 | -105.9787
27.0265 | -105.9787
27.0302 | -105.979
27.0339 | -105.9781
27.0377 | -105.977
27.0414 | -105.976
27.0451 | -105.9752
27.0489 | -105.9735
27.0526 | -105.9724
27.0564 | -105.9707
27.0601 | -105.9697
27.0638 | -105.968
27.0676 | -105.9663
27.0713 | -105.9665
27.075 | -105.9669
27.0788 | -105.9674
27.0825 | -105.9682
27.0862 | -105.9694
27.09 | -105.9702
27.0937 | -105.9716
27.0974 | -105.972
27.1012 | -105.9723
27.1049 | -105.9733
27.1086 | -105.9745
27.1124 | -105.9777
27.1161 | -105.9813
27.1198 | -105.984
27.1236 | -105.9878
27.1273 | -105.9911
27.1311 | -105.9956
27.1348 | -105.9988
27.1385 | -106.001
27.1423 | -106.0039
27.146 | -106.0089
27.1497 | -106.0156
27.1535 | -106.0228
27.1572 | -106.0305
27.1609 | -106.037
27.1647 | -106.0441
27.1684 | -106.0503
27.1721 | -106.058
27.1759 | -106.066
27.1796 | -106.0726
27.1833 | -106.0784
27.1871 | -106.0843
27.1908 | -106.0907
27.1945 | -106.0975
27.1983 | -106.1045
27.202 | -106.1093
27.2058 | -106.1162
27.2095 | -106.1245
27.2132 | -106.1308
27.217 | -106.1364
27.2207 | -106.142
27.2244 | -106.1492
27.2282 | -106.1566
27.2319 | -106.1615
27.2356 | -106.1676
27.2394 | -106.1734
27.2431 | -106.1784
27.2468 | -106.1824
27.2506 | -106.1874
27.2543 | -106.1922
27.258 | -106.1957
27.2618 | -106.1982
27.2655 | -106.2004
27.2692 | -106.2028
27.273 | -106.205
27.2767 | -106.2069
27.2805 | -106.2071
27.2842 | -106.2061
27.2879 | -106.2049
27.2917 | -106.2052
27.2954 | -106.2044
27.2991 | -106.202
27.3029 | -106.1991
27.3066 | -106.1961
27.3103 | -106.1929
27.3141 | -106.1899
27.3178 | -106.1868
27.3215 | -106.1833
27.3253 | -106.1798
27.329 | -106.1764
27.3327 | -106.1711
27.3365 | -106.1642
27.3402 | -106.1572
27.3439 | -106.1512
27.3477 | -106.1454
27.3514 | -106.1399
27.3552 | -106.1333
27.3589 | -106.127
27.3626 | -106.1212
27.3664 | -106.1143
27.3701 | -106.1079
27.3738 | -106.1016
27.3776 | -106.0942
27.3813 | -106.0845
27.385 | -106.0752
27.3888 | -106.066
27.3925 | -106.0578
27.3962 | -106.0507
27.4 | -106.044
27.4037 | -106.0374
27.4074 | -106.03
27.4112 | -106.0236
27.4149 | -106.0158
27.4186 | -106.0091
27.4224 | -106.0029
27.4261 | -105.9982
27.4299 | -105.9928
27.4336 | -105.9878
27.4373 | -105.9838
27.4411 | -105.9798
27.4448 | -105.9747
27.4485 | -105.9709
27.4523 | -105.9691
27.456 | -105.9653
27.4597 | -105.962
27.4635 | -105.9608
27.4672 | -105.9593
27.4709 | -105.9576
27.4747 | -105.9577
27.4784 | -105.9603
27.4821 | -105.9593
27.4859 | -105.959
27.4896 | -105.9587
27.4933 | -105.9584
27.4971 | -105.9593
27.5008 | -105.9612
27.5046 | -105.9625
27.5083 | -105.9638
27.512 | -105.9645
27.5158 | -105.9667
27.5195 | -105.9682
27.5232 | -105.9706
27.527 | -105.9741
27.5307 | -105.9769
27.5344 | -105.9783
27.5382 | -105.9807
27.5419 | -105.9833
27.5456 | -105.9866
27.5494 | -105.9897
27.5531 | -105.9922
27.5568 | -105.9956
27.5606 | -105.9978
27.5643 | -105.9989
27.568 | -106.0011
27.5718 | -106.0049
27.5755 | -106.0079
27.5793 | -106.0098
27.583 | -106.012
27.5867 | -106.0131
27.5905 | -106.0162
27.5942 | -106.0181
27.5979 | -106.0192
27.6017 | -106.021
27.6054 | -106.024
27.6091 | -106.026
27.6129 | -106.0272
27.6166 | -106.0269
27.6203 | -106.0273
27.6241 | -106.0281
27.6278 | -106.0297
27.6315 | -106.0298
27.6353 | -106.0293
27.639 | -106.0292
27.6427 | -106.0285
27.6465 | -106.0286
27.6502 | -106.03
27.654 | -106.0289
27.6577 | -106.0262
27.6614 | -106.023
27.6652 | -106.0217
27.6689 | -106.0213
27.6726 | -106.0198
27.6764 | -106.0158
27.6801 | -106.0108
27.6838 | -106.006
27.6876 | -106.0016
27.6913 | -105.9982
27.695 | -105.9954
27.6988 | -105.9918
27.7025 | -105.9868
27.7062 | -105.9803
27.71 | -105.9739
27.7137 | -105.9693
27.7174 | -105.9647
27.7212 | -105.9618
27.7249 | -105.9597
27.7287 | -105.9564
27.7324 | -105.9494
27.7361 | -105.9426
27.7399 | -105.936
27.7436 | -105.9302
27.7473 | -105.9244
27.7511 | -105.9172
27.7548 | -105.9108
27.7585 | -105.9054
27.7623 | -105.9013
27.766 | -105.897
27.7697 | -105.8931
27.7735 | -105.8889
27.7772 | -105.8838
27.7809 | -105.8791
27.7847 | -105.8741
27.7884 | -105.8699
27.7921 | -105.8655
27.7959 | -105.8614
27.7996 | -105.857
27.8034 | -105.8537
27.8071 | -105.8493
27.8108 | -105.845
27.8146 | -105.8416
27.8183 | -105.8382
27.822 | -105.8353
27.8258 | -105.8344
27.8295 | -105.8343
27.8332 | -105.8333
27.837 | -105.8318
27.8407 | -105.8297
27.8444 | -105.8289
27.8482 | -105.8287
27.8519 | -105.8268
27.8556 | -105.8243
27.8594 | -105.8234
27.8631 | -105.8234
27.8668 | -105.8239
27.8706 | -105.8236
27.8743 | -105.8246
27.8781 | -105.8264
27.8818 | -105.827
27.8855 | -105.8284
27.8893 | -105.8296
27.893 | -105.8295
27.8967 | -105.8299
27.9005 | -105.8297
27.9042 | -105.831
27.9079 | -105.8331
27.9117 | -105.8355
27.9154 | -105.8382
27.9191 | -105.8396
27.9229 | -105.8396
27.9266 | -105.8387
27.9303 | -105.8361
27.9341 | -105.8336
27.9378 | -105.8316
27.9415 | -105.8307
27.9453 | -105.8311
27.949 | -105.8299
27.9528 | -105.8295
27.9565 | -105.8298
27.9602 | -105.8292
27.964 | -105.8269
27.9677 | -105.8241
27.9714 | -105.8228
27.9752 | -105.8205
27.9789 | -105.8177
27.9826 | -105.8149
27.9864 | -105.8117
27.9901 | -105.8069
27.9938 | -105.8011
27.9976 | -105.7945
28.0013 | -105.7877