## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 278 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 284
Hood | 02-92B/P | 223
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1008.4 | hPa |  | PRC PRO 01
Humidity | 59.9 | %/Hr
Internal temperature of calibration bench | 24.31 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1018 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.8 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.07 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.01 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3082 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 23/08/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 28/08/2018
Technical Manager | Anthony HUE | Approval |  | 30/08/2018
CEO | Robert GLEMOT | Approval |  | 30/08/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3486 | 57.4431
0.012209 | -35.9071 | 38.2334
0.018313 | -34.8414 | 28.0909
0.024417 | -34.3835 | 21.8386
0.030521 | -34.2795 | 18.1037
0.036626 | -34.1388 | 15.0308
0.04273 | -34.0459 | 12.6392
0.048834 | -34.0023 | 11.138
0.054938 | -33.9358 | 10.0104
0.061043 | -33.8728 | 8.9163
0.067147 | -33.8648 | 8.0757
0.073251 | -33.8688 | 7.297
0.079355 | -33.8907 | 6.6244
0.08546 | -33.9071 | 6.0286
0.091564 | -33.9029 | 5.5971
0.097668 | -33.8845 | 5.3639
0.10377 | -33.8629 | 5.1166
0.10988 | -33.8462 | 4.8394
0.11598 | -33.8378 | 4.5445
0.12209 | -33.841 | 4.2595
0.12819 | -33.8528 | 3.9734
0.13429 | -33.8656 | 3.6808
0.1404 | -33.8692 | 3.3969
0.1465 | -33.8713 | 3.1345
0.15261 | -33.8683 | 2.9102
0.15871 | -33.8613 | 2.746
0.16482 | -33.853 | 2.6432
0.17092 | -33.8516 | 2.607
0.17702 | -33.8501 | 2.5358
0.18313 | -33.85 | 2.4587
0.18923 | -33.8503 | 2.3483
0.19534 | -33.8496 | 2.2188
0.20144 | -33.8487 | 2.072
0.20754 | -33.8461 | 1.9475
0.21365 | -33.8455 | 1.8302
0.21975 | -33.8456 | 1.7308
0.22586 | -33.8473 | 1.6433
0.23196 | -33.8505 | 1.5722
0.23807 | -33.8554 | 1.5112
0.24417 | -33.8602 | 1.4526
0.25027 | -33.8649 | 1.3901
0.25638 | -33.8679 | 1.3271
0.26248 | -33.8703 | 1.2586
0.26859 | -33.8718 | 1.1912
0.27469 | -33.8727 | 1.1173
0.2808 | -33.8743 | 1.042
0.2869 | -33.8757 | 0.9695
0.293 | -33.8772 | 0.9014
0.29911 | -33.8786 | 0.84138
0.30521 | -33.8801 | 0.78248
0.31132 | -33.8813 | 0.74588
0.31742 | -33.8831 | 0.70462
0.32353 | -33.8839 | 0.67146
0.32963 | -33.8852 | 0.64471
0.33573 | -33.8868 | 0.62669
0.34184 | -33.8887 | 0.60716
0.34794 | -33.8914 | 0.5839
0.35405 | -33.8941 | 0.56073
0.36015 | -33.8972 | 0.53594
0.36626 | -33.8997 | 0.50858
0.37236 | -33.9023 | 0.47125
0.37846 | -33.9049 | 0.42571
0.38457 | -33.9061 | 0.36677
0.39067 | -33.9074 | 0.31901
0.39678 | -33.9083 | 0.26577
0.40288 | -33.909 | 0.21023
0.40899 | -33.909 | 0.157
0.41509 | -33.9085 | 0.10515
0.42119 | -33.9087 | 0.061243
0.4273 | -33.9088 | 0.014065
0.4334 | -33.909 | -0.033216
0.43951 | -33.9095 | -0.077873
0.44561 | -33.9107 | -0.11871
0.45172 | -33.9127 | -0.1514
0.45782 | -33.9139 | -0.18959
0.46392 | -33.9155 | -0.22699
0.47003 | -33.9169 | -0.26035
0.47613 | -33.9182 | -0.29564
0.48224 | -33.919 | -0.33186
0.48834 | -33.9183 | -0.3684
0.49445 | -33.9183 | -0.40518
0.50055 | -33.9179 | -0.4387
0.50665 | -33.9174 | -0.47466
0.51276 | -33.9165 | -0.5105
0.51886 | -33.9156 | -0.54607
0.52497 | -33.9144 | -0.57118
0.53107 | -33.9138 | -0.60206
0.53717 | -33.9132 | -0.63138
0.54328 | -33.9125 | -0.65512
0.54938 | -33.9125 | -0.67828
0.55549 | -33.9127 | -0.69974
0.56159 | -33.914 | -0.7109
0.5677 | -33.9145 | -0.72534
0.5738 | -33.9158 | -0.74167
0.5799 | -33.9168 | -0.75642
0.58601 | -33.9182 | -0.76586
0.59211 | -33.9201 | -0.7711
0.59822 | -33.9217 | -0.77978
0.60432 | -33.9234 | -0.79143
0.61043 | -33.9249 | -0.79842
0.61653 | -33.9264 | -0.80788
0.62263 | -33.9282 | -0.81628
0.62874 | -33.929 | -0.82598
0.63484 | -33.93 | -0.83422
0.64095 | -33.9307 | -0.84342
0.64705 | -33.9312 | -0.85221
0.65316 | -33.9315 | -0.86401
0.65926 | -33.9317 | -0.8795
0.66536 | -33.9307 | -0.89689
0.67147 | -33.9304 | -0.90959
0.67757 | -33.9301 | -0.92356
0.68368 | -33.9295 | -0.94256
0.68978 | -33.9288 | -0.95963
0.69589 | -33.928 | -0.97875
0.70199 | -33.9271 | -1.0021
0.70809 | -33.9264 | -1.0214
0.7142 | -33.9258 | -1.0421
0.7203 | -33.9248 | -1.0616
0.72641 | -33.9243 | -1.0847
0.73251 | -33.9238 | -1.1051
0.73862 | -33.9243 | -1.1273
0.74472 | -33.9242 | -1.1477
0.75082 | -33.9241 | -1.1704
0.75693 | -33.9247 | -1.1928
0.76303 | -33.9255 | -1.2153
0.76914 | -33.9277 | -1.2361
0.77524 | -33.9289 | -1.256
0.78135 | -33.9305 | -1.2765
0.78745 | -33.9322 | -1.2967
0.79355 | -33.9337 | -1.3169
0.79966 | -33.9358 | -1.3366
0.80576 | -33.9384 | -1.3566
0.81187 | -33.9403 | -1.3787
0.81797 | -33.9425 | -1.4005
0.82408 | -33.9445 | -1.4214
0.83018 | -33.9465 | -1.4424
0.83628 | -33.9487 | -1.4613
0.84239 | -33.9501 | -1.4824
0.84849 | -33.9518 | -1.5047
0.8546 | -33.9535 | -1.5249
0.8607 | -33.955 | -1.5462
0.86681 | -33.9563 | -1.5656
0.87291 | -33.9575 | -1.5848
0.87901 | -33.9581 | -1.6033
0.88512 | -33.959 | -1.6219
0.89122 | -33.9599 | -1.6405
0.89733 | -33.9605 | -1.6563
0.90343 | -33.9611 | -1.6743
0.90953 | -33.9614 | -1.6886
0.91564 | -33.9616 | -1.7046
0.92174 | -33.962 | -1.7197
0.92785 | -33.9622 | -1.7343
0.93395 | -33.9625 | -1.7481
0.94006 | -33.9628 | -1.7607
0.94616 | -33.9626 | -1.7722
0.95226 | -33.9629 | -1.7847
0.95837 | -33.9632 | -1.7984
0.96447 | -33.9635 | -1.811
0.97058 | -33.9635 | -1.8241
0.97668 | -33.9634 | -1.8372
0.98279 | -33.9632 | -1.8496
0.98889 | -33.9632 | -1.8598
0.99499 | -33.9632 | -1.873
1.0011 | -33.9633 | -1.8832
1.0072 | -33.963 | -1.8945
1.0133 | -33.9628 | -1.9047
1.0194 | -33.9624 | -1.9175
1.0255 | -33.9622 | -1.931
1.0316 | -33.9617 | -1.9426
1.0377 | -33.9613 | -1.9561
1.0438 | -33.9608 | -1.9685
1.0499 | -33.9601 | -1.9839
1.056 | -33.9594 | -1.9981
1.0621 | -33.9589 | -2.0123
1.0682 | -33.9585 | -2.0246
1.0743 | -33.9578 | -2.0401
1.0805 | -33.9574 | -2.0545
1.0866 | -33.9565 | -2.0698
1.0927 | -33.9556 | -2.0843
1.0988 | -33.9551 | -2.0992
1.1049 | -33.9543 | -2.1134
1.111 | -33.9535 | -2.129
1.1171 | -33.9528 | -2.1436
1.1232 | -33.9524 | -2.159
1.1293 | -33.952 | -2.1731
1.1354 | -33.9518 | -2.1884
1.1415 | -33.9514 | -2.2041
1.1476 | -33.9509 | -2.2198
1.1537 | -33.9504 | -2.2347
1.1598 | -33.9505 | -2.2498
1.1659 | -33.9504 | -2.2633
1.172 | -33.9501 | -2.2773
1.1781 | -33.9499 | -2.2914
1.1842 | -33.9498 | -2.3058
1.1903 | -33.9503 | -2.3166
1.1964 | -33.9502 | -2.3301
1.2025 | -33.9504 | -2.3442
1.2086 | -33.9508 | -2.3568
1.2147 | -33.9508 | -2.3693
1.2209 | -33.9513 | -2.3805
1.227 | -33.9521 | -2.391
1.2331 | -33.9526 | -2.403
1.2392 | -33.9531 | -2.4148
1.2453 | -33.9536 | -2.4257
1.2514 | -33.9541 | -2.4354
1.2575 | -33.9548 | -2.4462
1.2636 | -33.9557 | -2.4546
1.2697 | -33.9564 | -2.4635
1.2758 | -33.9572 | -2.4744
1.2819 | -33.9581 | -2.484
1.288 | -33.9588 | -2.4946
1.2941 | -33.9598 | -2.5026
1.3002 | -33.9606 | -2.5101
1.3063 | -33.9613 | -2.5187
1.3124 | -33.962 | -2.5273
1.3185 | -33.9629 | -2.5359
1.3246 | -33.9637 | -2.5445
1.3307 | -33.9647 | -2.5525
1.3368 | -33.9656 | -2.5608
1.3429 | -33.9664 | -2.5692
1.349 | -33.967 | -2.578
1.3551 | -33.9678 | -2.5869
1.3613 | -33.9685 | -2.5946
1.3674 | -33.9693 | -2.6043
1.3735 | -33.9698 | -2.6136
1.3796 | -33.9704 | -2.6222
1.3857 | -33.9708 | -2.631
1.3918 | -33.9715 | -2.6409
1.3979 | -33.9722 | -2.6497
1.404 | -33.9728 | -2.6609
1.4101 | -33.9733 | -2.6711
1.4162 | -33.9738 | -2.6814
1.4223 | -33.9744 | -2.6926
1.4284 | -33.975 | -2.7024
1.4345 | -33.9753 | -2.7137
1.4406 | -33.9757 | -2.7274
1.4467 | -33.9763 | -2.7389
1.4528 | -33.9768 | -2.7502
1.4589 | -33.9774 | -2.762
1.465 | -33.9777 | -2.7728
1.4711 | -33.9783 | -2.7841
1.4772 | -33.979 | -2.7976
1.4833 | -33.9795 | -2.8107
1.4894 | -33.9802 | -2.8229
1.4955 | -33.9808 | -2.8349
1.5016 | -33.9815 | -2.8475
1.5078 | -33.982 | -2.8617
1.5139 | -33.9825 | -2.8748
1.52 | -33.983 | -2.8876
1.5261 | -33.9834 | -2.9003
1.5322 | -33.9838 | -2.9145
1.5383 | -33.9843 | -2.9272
1.5444 | -33.985 | -2.94
1.5505 | -33.9856 | -2.9536
1.5566 | -33.9862 | -2.9668
1.5627 | -33.9868 | -2.9808
1.5688 | -33.9875 | -2.9946
1.5749 | -33.9879 | -3.0074
1.581 | -33.9883 | -3.0211
1.5871 | -33.9888 | -3.0351
1.5932 | -33.9893 | -3.0481
1.5993 | -33.9898 | -3.0626
1.6054 | -33.9903 | -3.0761
1.6115 | -33.9908 | -3.0905
1.6176 | -33.9912 | -3.1032
1.6237 | -33.9915 | -3.1171
1.6298 | -33.9918 | -3.1298
1.6359 | -33.9923 | -3.1428
1.642 | -33.9928 | -3.1549
1.6482 | -33.9931 | -3.1681
1.6543 | -33.9936 | -3.182
1.6604 | -33.994 | -3.1941
1.6665 | -33.9941 | -3.2068
1.6726 | -33.9945 | -3.2194
1.6787 | -33.9949 | -3.2323
1.6848 | -33.9949 | -3.245
1.6909 | -33.9952 | -3.258
1.697 | -33.9955 | -3.2703
1.7031 | -33.9957 | -3.2832
1.7092 | -33.9957 | -3.2955
1.7153 | -33.996 | -3.3076
1.7214 | -33.9958 | -3.3199
1.7275 | -33.9958 | -3.3326
1.7336 | -33.9958 | -3.3457
1.7397 | -33.996 | -3.3581
1.7458 | -33.9961 | -3.3704
1.7519 | -33.9962 | -3.3827
1.758 | -33.9963 | -3.3947
1.7641 | -33.9965 | -3.4064
1.7702 | -33.9965 | -3.4195
1.7763 | -33.9966 | -3.4311
1.7824 | -33.9966 | -3.4436
1.7885 | -33.9967 | -3.4562
1.7947 | -33.9969 | -3.4682
1.8008 | -33.9969 | -3.4795
1.8069 | -33.997 | -3.4911
1.813 | -33.9972 | -3.5034
1.8191 | -33.9972 | -3.5162
1.8252 | -33.9974 | -3.529
1.8313 | -33.9976 | -3.5413
1.8374 | -33.9978 | -3.5532
1.8435 | -33.9979 | -3.5651
1.8496 | -33.9981 | -3.5783
1.8557 | -33.9983 | -3.5911
1.8618 | -33.9987 | -3.6032
1.8679 | -33.9989 | -3.6153
1.874 | -33.9991 | -3.6272
1.8801 | -33.9993 | -3.6397
1.8862 | -33.9997 | -3.6516
1.8923 | -33.9999 | -3.6638
1.8984 | -34.0001 | -3.6758
1.9045 | -34.0002 | -3.6873
1.9106 | -34.0004 | -3.6994
1.9167 | -34.0005 | -3.7116
1.9228 | -34.0007 | -3.7234
1.9289 | -34.0009 | -3.7352
1.9351 | -34.0012 | -3.7459
1.9412 | -34.0015 | -3.757
1.9473 | -34.0018 | -3.769
1.9534 | -34.0021 | -3.7804
1.9595 | -34.0024 | -3.7921
1.9656 | -34.0027 | -3.8043
1.9717 | -34.003 | -3.8159
1.9778 | -34.0035 | -3.8272
1.9839 | -34.0038 | -3.8386
1.99 | -34.004 | -3.8495
1.9961 | -34.0042 | -3.8616
2.0022 | -34.0045 | -3.873
2.0083 | -34.0047 | -3.8841
2.0144 | -34.005 | -3.8954
2.0205 | -34.0054 | -3.907
2.0266 | -34.0058 | -3.9183
2.0327 | -34.0061 | -3.9298
2.0388 | -34.0066 | -3.9413
2.0449 | -34.0069 | -3.9526
2.051 | -34.0072 | -3.9639
2.0571 | -34.0076 | -3.9752
2.0632 | -34.0079 | -3.9866
2.0693 | -34.0081 | -3.9988
2.0754 | -34.0085 | -4.0101
2.0816 | -34.0088 | -4.0222
2.0877 | -34.0091 | -4.0337
2.0938 | -34.0095 | -4.0458
2.0999 | -34.0098 | -4.0575
2.106 | -34.01 | -4.0691
2.1121 | -34.0102 | -4.0811
2.1182 | -34.0106 | -4.0918
2.1243 | -34.0109 | -4.1025
2.1304 | -34.0111 | -4.1135
2.1365 | -34.0115 | -4.1253
2.1426 | -34.0117 | -4.1373
2.1487 | -34.0121 | -4.1485
2.1548 | -34.0126 | -4.1603
2.1609 | -34.0131 | -4.1718
2.167 | -34.0135 | -4.1829
2.1731 | -34.0139 | -4.195
2.1792 | -34.0141 | -4.206
2.1853 | -34.0146 | -4.2179
2.1914 | -34.0149 | -4.2293
2.1975 | -34.0154 | -4.2408
2.2036 | -34.0156 | -4.2529
2.2097 | -34.0159 | -4.2643
2.2158 | -34.0164 | -4.2748
2.222 | -34.0167 | -4.2852
2.2281 | -34.0169 | -4.2966
2.2342 | -34.0173 | -4.3079
2.2403 | -34.0175 | -4.3185
2.2464 | -34.0178 | -4.3292
2.2525 | -34.0181 | -4.3406
2.2586 | -34.0185 | -4.3511
2.2647 | -34.0189 | -4.3614
2.2708 | -34.0192 | -4.3726
2.2769 | -34.0196 | -4.3832
2.283 | -34.0198 | -4.3943
2.2891 | -34.0201 | -4.4052
2.2952 | -34.0203 | -4.4157
2.3013 | -34.0206 | -4.4265
2.3074 | -34.0209 | -4.4368
2.3135 | -34.0211 | -4.4482
2.3196 | -34.0214 | -4.4583
2.3257 | -34.0216 | -4.4688
2.3318 | -34.0218 | -4.4798
2.3379 | -34.022 | -4.4905
2.344 | -34.0222 | -4.5018
2.3501 | -34.0224 | -4.5125
2.3562 | -34.0226 | -4.5228
2.3623 | -34.0228 | -4.5331
2.3685 | -34.023 | -4.5435
2.3746 | -34.0231 | -4.5548
2.3807 | -34.0234 | -4.5651
2.3868 | -34.0236 | -4.5758
2.3929 | -34.0238 | -4.5864
2.399 | -34.0239 | -4.5967
2.4051 | -34.0243 | -4.6071
2.4112 | -34.0245 | -4.6177
2.4173 | -34.0247 | -4.6282
2.4234 | -34.0248 | -4.639
2.4295 | -34.025 | -4.6498
2.4356 | -34.0252 | -4.6605
2.4417 | -34.0254 | -4.6716
2.4478 | -34.0255 | -4.6825
2.4539 | -34.0257 | -4.6926
2.46 | -34.0259 | -4.7041
2.4661 | -34.0261 | -4.7151
2.4722 | -34.0263 | -4.7252
2.4783 | -34.0264 | -4.7356
2.4844 | -34.0265 | -4.7466
2.4905 | -34.0267 | -4.7575
2.4966 | -34.0269 | -4.7681
2.5027 | -34.027 | -4.7786
2.5089 | -34.0272 | -4.7894
2.515 | -34.0273 | -4.8001
2.5211 | -34.0274 | -4.8108
2.5272 | -34.0275 | -4.8217
2.5333 | -34.0277 | -4.8326
2.5394 | -34.0279 | -4.8434
2.5455 | -34.0281 | -4.8547
2.5516 | -34.0283 | -4.865
2.5577 | -34.0286 | -4.8764
2.5638 | -34.0288 | -4.8882
2.5699 | -34.0288 | -4.8996
2.576 | -34.0289 | -4.9113
2.5821 | -34.0291 | -4.9225
2.5882 | -34.0293 | -4.9337
2.5943 | -34.0294 | -4.9445
2.6004 | -34.0297 | -4.9554
2.6065 | -34.0299 | -4.9657
2.6126 | -34.03 | -4.9762
2.6187 | -34.0302 | -4.9866
2.6248 | -34.0304 | -4.9974
2.6309 | -34.0306 | -5.0089
2.637 | -34.0309 | -5.0206
2.6431 | -34.0311 | -5.031
2.6492 | -34.0313 | -5.0424
2.6554 | -34.0315 | -5.0532
2.6615 | -34.0317 | -5.0643
2.6676 | -34.032 | -5.0757
2.6737 | -34.0323 | -5.0869
2.6798 | -34.0327 | -5.0981
2.6859 | -34.0329 | -5.1094
2.692 | -34.0331 | -5.1203
2.6981 | -34.0334 | -5.1314
2.7042 | -34.0337 | -5.1426
2.7103 | -34.0339 | -5.1532
2.7164 | -34.0343 | -5.1642
2.7225 | -34.0346 | -5.1756
2.7286 | -34.0349 | -5.1866
2.7347 | -34.0351 | -5.1971
2.7408 | -34.0355 | -5.2071
2.7469 | -34.0358 | -5.2184
2.753 | -34.0361 | -5.229
2.7591 | -34.0365 | -5.2404
2.7652 | -34.0368 | -5.2519
2.7713 | -34.037 | -5.2627
2.7774 | -34.0373 | -5.2737
2.7835 | -34.0375 | -5.2847
2.7896 | -34.0378 | -5.2951
2.7958 | -34.0381 | -5.3059
2.8019 | -34.0385 | -5.3163
2.808 | -34.0388 | -5.3271
2.8141 | -34.0392 | -5.3383
2.8202 | -34.0396 | -5.3492
2.8263 | -34.0399 | -5.3601
2.8324 | -34.0403 | -5.3715
2.8385 | -34.0408 | -5.3824
2.8446 | -34.0412 | -5.3931
2.8507 | -34.0417 | -5.4039
2.8568 | -34.0421 | -5.4144
2.8629 | -34.0424 | -5.4249
2.869 | -34.0428 | -5.436
2.8751 | -34.0431 | -5.4468
2.8812 | -34.0436 | -5.4581
2.8873 | -34.044 | -5.469
2.8934 | -34.0443 | -5.48
2.8995 | -34.0447 | -5.4912
2.9056 | -34.0451 | -5.5023
2.9117 | -34.0455 | -5.5133
2.9178 | -34.0459 | -5.5241
2.9239 | -34.0463 | -5.5357
2.93 | -34.0468 | -5.5472
2.9361 | -34.0471 | -5.5579
2.9423 | -34.0475 | -5.5686
2.9484 | -34.048 | -5.5797
2.9545 | -34.0484 | -5.5907
2.9606 | -34.0489 | -5.6021
2.9667 | -34.0493 | -5.6132
2.9728 | -34.0497 | -5.6245
2.9789 | -34.0502 | -5.6356
2.985 | -34.0507 | -5.6469
2.9911 | -34.0511 | -5.6581
2.9972 | -34.0516 | -5.6693
3.0033 | -34.0521 | -5.6808
3.0094 | -34.0525 | -5.6919
3.0155 | -34.053 | -5.7027
3.0216 | -34.0534 | -5.714
3.0277 | -34.0538 | -5.7257
3.0338 | -34.0542 | -5.7371
3.0399 | -34.0547 | -5.7485
3.046 | -34.0551 | -5.7601
3.0521 | -34.0554 | -5.7711
3.0582 | -34.0558 | -5.7832
3.0643 | -34.0562 | -5.7946
3.0704 | -34.0565 | -5.8065
3.0765 | -34.0568 | -5.8176
3.0827 | -34.0572 | -5.8288
3.0888 | -34.0576 | -5.84
3.0949 | -34.058 | -5.8513
3.101 | -34.0584 | -5.8628
3.1071 | -34.0588 | -5.8745
3.1132 | -34.0592 | -5.8863
3.1193 | -34.0596 | -5.8978
3.1254 | -34.06 | -5.9091
3.1315 | -34.0604 | -5.9211
3.1376 | -34.0607 | -5.9324
3.1437 | -34.0612 | -5.9441
3.1498 | -34.0615 | -5.9562
3.1559 | -34.0619 | -5.9676
3.162 | -34.0622 | -5.9797
3.1681 | -34.0625 | -5.9916
3.1742 | -34.0628 | -6.0032
3.1803 | -34.0632 | -6.0149
3.1864 | -34.0636 | -6.0261
3.1925 | -34.0639 | -6.0376
3.1986 | -34.0642 | -6.0496
3.2047 | -34.0645 | -6.0617
3.2108 | -34.0648 | -6.0727
3.2169 | -34.0652 | -6.0844
3.223 | -34.0654 | -6.0961
3.2292 | -34.0657 | -6.108
3.2353 | -34.066 | -6.1187
3.2414 | -34.0662 | -6.1301
3.2475 | -34.0665 | -6.1415
3.2536 | -34.0668 | -6.1529
3.2597 | -34.0671 | -6.1638
3.2658 | -34.0674 | -6.1748
3.2719 | -34.0676 | -6.1858
3.278 | -34.0679 | -6.1973
3.2841 | -34.0682 | -6.2084
3.2902 | -34.0685 | -6.2198
3.2963 | -34.0688 | -6.2313
3.3024 | -34.0689 | -6.2423
3.3085 | -34.0692 | -6.2539
3.3146 | -34.0694 | -6.265
3.3207 | -34.0697 | -6.2765
3.3268 | -34.0699 | -6.288
3.3329 | -34.0702 | -6.2986
3.339 | -34.0704 | -6.3094
3.3451 | -34.0706 | -6.3205
3.3512 | -34.0708 | -6.3312
3.3573 | -34.0711 | -6.3424
3.3634 | -34.0713 | -6.3535
3.3696 | -34.0715 | -6.3648
3.3757 | -34.0717 | -6.3755
3.3818 | -34.072 | -6.3864
3.3879 | -34.0722 | -6.3974
3.394 | -34.0723 | -6.4088
3.4001 | -34.0726 | -6.4196
3.4062 | -34.0727 | -6.4305
3.4123 | -34.0729 | -6.441
3.4184 | -34.0731 | -6.4523
3.4245 | -34.0734 | -6.463
3.4306 | -34.0735 | -6.4737
3.4367 | -34.0737 | -6.4842
3.4428 | -34.0739 | -6.4945
3.4489 | -34.0741 | -6.5054
3.455 | -34.0743 | -6.5162
3.4611 | -34.0745 | -6.5275
3.4672 | -34.0747 | -6.5386
3.4733 | -34.0749 | -6.5496
3.4794 | -34.0751 | -6.5605
3.4855 | -34.0753 | -6.5708
3.4916 | -34.0755 | -6.5811
3.4977 | -34.0757 | -6.5925
3.5038 | -34.0759 | -6.6033
3.5099 | -34.0761 | -6.6135
3.5161 | -34.0763 | -6.6239
3.5222 | -34.0764 | -6.6342
3.5283 | -34.0766 | -6.645
3.5344 | -34.0767 | -6.6558
3.5405 | -34.0769 | -6.6663
3.5466 | -34.0772 | -6.6773
3.5527 | -34.0773 | -6.6881
3.5588 | -34.0774 | -6.6987
3.5649 | -34.0775 | -6.7094
3.571 | -34.0777 | -6.72
3.5771 | -34.0779 | -6.7311
3.5832 | -34.0781 | -6.741
3.5893 | -34.0783 | -6.7516
3.5954 | -34.0785 | -6.7624
3.6015 | -34.0787 | -6.7733
3.6076 | -34.079 | -6.7838
3.6137 | -34.0792 | -6.7939
3.6198 | -34.0795 | -6.8041
3.6259 | -34.0796 | -6.8151
3.632 | -34.0799 | -6.8257
3.6381 | -34.0801 | -6.8365
3.6442 | -34.0802 | -6.8471
3.6503 | -34.0805 | -6.8572
3.6565 | -34.0808 | -6.868
3.6626 | -34.0809 | -6.8786
3.6687 | -34.0812 | -6.8893
3.6748 | -34.0815 | -6.8996
3.6809 | -34.0818 | -6.9103
3.687 | -34.082 | -6.9213
3.6931 | -34.0823 | -6.9315
3.6992 | -34.0825 | -6.9423
3.7053 | -34.0828 | -6.9528
3.7114 | -34.0829 | -6.9635
3.7175 | -34.0832 | -6.9742
3.7236 | -34.0835 | -6.9843
3.7297 | -34.0838 | -6.9951
3.7358 | -34.084 | -7.006
3.7419 | -34.0842 | -7.0169
3.748 | -34.0845 | -7.0274
3.7541 | -34.0847 | -7.0381
3.7602 | -34.0851 | -7.0486
3.7663 | -34.0853 | -7.0591
3.7724 | -34.0856 | -7.0698
3.7785 | -34.0859 | -7.0802
3.7846 | -34.0861 | -7.0907
3.7907 | -34.0865 | -7.1017
3.7969 | -34.0868 | -7.1124
3.803 | -34.087 | -7.1226
3.8091 | -34.0873 | -7.1332
3.8152 | -34.0876 | -7.1438
3.8213 | -34.0879 | -7.1547
3.8274 | -34.0882 | -7.165
3.8335 | -34.0886 | -7.1758
3.8396 | -34.0889 | -7.1865
3.8457 | -34.0892 | -7.1973
3.8518 | -34.0896 | -7.2078
3.8579 | -34.09 | -7.2185
3.864 | -34.0903 | -7.2293
3.8701 | -34.0906 | -7.2404
3.8762 | -34.0909 | -7.2508
3.8823 | -34.0912 | -7.2617
3.8884 | -34.0915 | -7.2724
3.8945 | -34.0919 | -7.2833
3.9006 | -34.0922 | -7.2944
3.9067 | -34.0925 | -7.3056
3.9128 | -34.0928 | -7.3164
3.9189 | -34.0932 | -7.3269
3.925 | -34.0934 | -7.3377
3.9311 | -34.0937 | -7.348
3.9372 | -34.0941 | -7.3589
3.9434 | -34.0944 | -7.3696
3.9495 | -34.0948 | -7.3805
3.9556 | -34.0952 | -7.391
3.9617 | -34.0955 | -7.4013
3.9678 | -34.0959 | -7.4117
3.9739 | -34.0961 | -7.4224
3.98 | -34.0964 | -7.4332
3.9861 | -34.0967 | -7.4439
3.9922 | -34.097 | -7.4546
3.9983 | -34.0973 | -7.4649
4.0044 | -34.0976 | -7.4751
4.0105 | -34.0979 | -7.4858
4.0166 | -34.0983 | -7.4966
4.0227 | -34.0986 | -7.5073
4.0288 | -34.0989 | -7.5184
4.0349 | -34.0991 | -7.5288
4.041 | -34.0994 | -7.5395
4.0471 | -34.0997 | -7.5503
4.0532 | -34.1 | -7.5605
4.0593 | -34.1003 | -7.5714
4.0654 | -34.1006 | -7.5819
4.0715 | -34.101 | -7.593
4.0776 | -34.1013 | -7.6036
4.0838 | -34.1016 | -7.6148
4.0899 | -34.1019 | -7.6252
4.096 | -34.1022 | -7.6358
4.1021 | -34.1025 | -7.6462
4.1082 | -34.1028 | -7.6566
4.1143 | -34.103 | -7.6669
4.1204 | -34.1034 | -7.6772
4.1265 | -34.1037 | -7.6875
4.1326 | -34.104 | -7.6979
4.1387 | -34.1043 | -7.7084
4.1448 | -34.1046 | -7.7191
4.1509 | -34.1048 | -7.7297
4.157 | -34.1051 | -7.7407
4.1631 | -34.1054 | -7.7515
4.1692 | -34.1057 | -7.762
4.1753 | -34.106 | -7.7725
4.1814 | -34.1062 | -7.7827
4.1875 | -34.1065 | -7.7936
4.1936 | -34.1068 | -7.8042
4.1997 | -34.1071 | -7.8144
4.2058 | -34.1073 | -7.8252
4.2119 | -34.1076 | -7.8356
4.218 | -34.1078 | -7.8458
4.2241 | -34.1081 | -7.8562
4.2303 | -34.1084 | -7.8665
4.2364 | -34.1087 | -7.877
4.2425 | -34.1089 | -7.8876
4.2486 | -34.1091 | -7.8982
4.2547 | -34.1092 | -7.9085
4.2608 | -34.1095 | -7.9191
4.2669 | -34.1097 | -7.9292
4.273 | -34.11 | -7.9399
4.2791 | -34.1102 | -7.9503
4.2852 | -34.1105 | -7.961
4.2913 | -34.1106 | -7.9711
4.2974 | -34.1109 | -7.9819
4.3035 | -34.1112 | -7.9927
4.3096 | -34.1114 | -8.0034
4.3157 | -34.1116 | -8.0136
4.3218 | -34.1118 | -8.0239
4.3279 | -34.1121 | -8.0344
4.334 | -34.1123 | -8.0447
4.3401 | -34.1125 | -8.0547
4.3462 | -34.1128 | -8.0653
4.3523 | -34.1131 | -8.0753
4.3584 | -34.1133 | -8.0853
4.3645 | -34.1135 | -8.0956
4.3707 | -34.1137 | -8.1057
4.3768 | -34.114 | -8.116
4.3829 | -34.1142 | -8.1267
4.389 | -34.1144 | -8.137
4.3951 | -34.1147 | -8.1472
4.4012 | -34.1149 | -8.1578
4.4073 | -34.1151 | -8.168
4.4134 | -34.1153 | -8.1782
4.4195 | -34.1156 | -8.1885
4.4256 | -34.1159 | -8.1987
4.4317 | -34.1161 | -8.2089
4.4378 | -34.1163 | -8.2188
4.4439 | -34.1165 | -8.2288
4.45 | -34.1167 | -8.239
4.4561 | -34.1169 | -8.2493
4.4622 | -34.1171 | -8.2596
4.4683 | -34.1173 | -8.2698
4.4744 | -34.1175 | -8.2804
4.4805 | -34.1178 | -8.2902
4.4866 | -34.118 | -8.3005
4.4927 | -34.1183 | -8.311
4.4988 | -34.1185 | -8.3211
4.5049 | -34.1188 | -8.331
4.511 | -34.119 | -8.3409
4.5172 | -34.1193 | -8.3511
4.5233 | -34.1195 | -8.3618
4.5294 | -34.1197 | -8.3718
4.5355 | -34.12 | -8.3821
4.5416 | -34.1202 | -8.3923
4.5477 | -34.1205 | -8.4025
4.5538 | -34.1207 | -8.4123
4.5599 | -34.121 | -8.4224
4.566 | -34.1212 | -8.4326
4.5721 | -34.1214 | -8.4428
4.5782 | -34.1217 | -8.4533
4.5843 | -34.122 | -8.4636
4.5904 | -34.1223 | -8.4739
4.5965 | -34.1225 | -8.4839
4.6026 | -34.1228 | -8.4945
4.6087 | -34.123 | -8.5046
4.6148 | -34.1232 | -8.5152
4.6209 | -34.1235 | -8.5253
4.627 | -34.1238 | -8.5357
4.6331 | -34.124 | -8.5458
4.6392 | -34.1242 | -8.5561
4.6453 | -34.1245 | -8.5663
4.6514 | -34.1248 | -8.5766
4.6576 | -34.125 | -8.5872
4.6637 | -34.1252 | -8.5971
4.6698 | -34.1254 | -8.6071
4.6759 | -34.1257 | -8.6177
4.682 | -34.1259 | -8.6279
4.6881 | -34.1261 | -8.6381
4.6942 | -34.1264 | -8.6482
4.7003 | -34.1266 | -8.6585
4.7064 | -34.1269 | -8.6686
4.7125 | -34.1272 | -8.6792
4.7186 | -34.1274 | -8.6892
4.7247 | -34.1277 | -8.6995
4.7308 | -34.1279 | -8.7097
4.7369 | -34.1282 | -8.7201
4.743 | -34.1284 | -8.7308
4.7491 | -34.1287 | -8.7414
4.7552 | -34.1289 | -8.7517
4.7613 | -34.1292 | -8.7621
4.7674 | -34.1294 | -8.7722
4.7735 | -34.1296 | -8.783
4.7796 | -34.13 | -8.7936
4.7857 | -34.1302 | -8.804
4.7918 | -34.1305 | -8.8141
4.7979 | -34.1308 | -8.8242
4.8041 | -34.1311 | -8.8346
4.8102 | -34.1313 | -8.8446
4.8163 | -34.1316 | -8.8548
4.8224 | -34.1319 | -8.8651
4.8285 | -34.1322 | -8.8751
4.8346 | -34.1325 | -8.8852
4.8407 | -34.1327 | -8.8956
4.8468 | -34.133 | -8.9062
4.8529 | -34.1333 | -8.9168
4.859 | -34.1335 | -8.9275
4.8651 | -34.1338 | -8.9376
4.8712 | -34.1341 | -8.9477
4.8773 | -34.1343 | -8.9581
4.8834 | -34.1346 | -8.9684
4.8895 | -34.1348 | -8.9785
4.8956 | -34.1351 | -8.9888
4.9017 | -34.1353 | -8.999
4.9078 | -34.1356 | -9.0097
4.9139 | -34.1359 | -9.0196
4.92 | -34.1362 | -9.0296
4.9261 | -34.1365 | -9.0396
4.9322 | -34.1368 | -9.0499
4.9383 | -34.1371 | -9.0602
4.9445 | -34.1374 | -9.0705
4.9506 | -34.1376 | -9.0806
4.9567 | -34.1379 | -9.0907
4.9628 | -34.1382 | -9.101
4.9689 | -34.1385 | -9.1118
4.975 | -34.1387 | -9.1224
4.9811 | -34.139 | -9.1323
4.9872 | -34.1394 | -9.1425
4.9933 | -34.1397 | -9.1532
4.9994 | -34.14 | -9.1639
5.0055 | -34.1403 | -9.1743
5.0116 | -34.1406 | -9.1851
5.0177 | -34.141 | -9.1957
5.0238 | -34.1413 | -9.2064
5.0299 | -34.1417 | -9.2171
5.036 | -34.142 | -9.2274
5.0421 | -34.1422 | -9.2377
5.0482 | -34.1425 | -9.248
5.0543 | -34.1428 | -9.2581
5.0604 | -34.1432 | -9.2685
5.0665 | -34.1435 | -9.2786
5.0726 | -34.1438 | -9.2893
5.0787 | -34.1441 | -9.2999
5.0848 | -34.1445 | -9.3104
5.091 | -34.1448 | -9.3206
5.0971 | -34.1451 | -9.3306
5.1032 | -34.1454 | -9.3414
5.1093 | -34.1457 | -9.3518
5.1154 | -34.146 | -9.3619
5.1215 | -34.1463 | -9.3721
5.1276 | -34.1466 | -9.3822
5.1337 | -34.147 | -9.3924
5.1398 | -34.1473 | -9.4025
5.1459 | -34.1476 | -9.4131
5.152 | -34.1479 | -9.4237
5.1581 | -34.1482 | -9.4342
5.1642 | -34.1486 | -9.4447
5.1703 | -34.1488 | -9.455
5.1764 | -34.1492 | -9.4652
5.1825 | -34.1495 | -9.4753
5.1886 | -34.1498 | -9.4853
5.1947 | -34.1501 | -9.4957
5.2008 | -34.1505 | -9.5061
5.2069 | -34.1508 | -9.5167
5.213 | -34.1511 | -9.5268
5.2191 | -34.1515 | -9.5373
5.2252 | -34.1518 | -9.5477
5.2314 | -34.1521 | -9.5583
5.2375 | -34.1524 | -9.5684
5.2436 | -34.1527 | -9.5788
5.2497 | -34.153 | -9.5888
5.2558 | -34.1533 | -9.5996
5.2619 | -34.1537 | -9.6103
5.268 | -34.154 | -9.6208
5.2741 | -34.1544 | -9.6313
5.2802 | -34.1547 | -9.642
5.2863 | -34.1551 | -9.6523
5.2924 | -34.1554 | -9.6629
5.2985 | -34.1558 | -9.6734
5.3046 | -34.1561 | -9.6839
5.3107 | -34.1564 | -9.6945
5.3168 | -34.1567 | -9.7049
5.3229 | -34.157 | -9.7154
5.329 | -34.1574 | -9.7258
5.3351 | -34.1577 | -9.7359
5.3412 | -34.158 | -9.7462
5.3473 | -34.1583 | -9.7567
5.3534 | -34.1587 | -9.7672
5.3595 | -34.159 | -9.7777
5.3656 | -34.1594 | -9.788
5.3717 | -34.1597 | -9.7988
5.3779 | -34.16 | -9.8096
5.384 | -34.1603 | -9.82
5.3901 | -34.1607 | -9.8307
5.3962 | -34.161 | -9.8409
5.4023 | -34.1614 | -9.8516
5.4084 | -34.1618 | -9.8619
5.4145 | -34.1621 | -9.8721
5.4206 | -34.1624 | -9.8826
5.4267 | -34.1627 | -9.8934
5.4328 | -34.1631 | -9.9039
5.4389 | -34.1634 | -9.914
5.445 | -34.1637 | -9.9248
5.4511 | -34.1641 | -9.935
5.4572 | -34.1644 | -9.9456
5.4633 | -34.1647 | -9.956
5.4694 | -34.165 | -9.9665
5.4755 | -34.1654 | -9.9766
5.4816 | -34.1657 | -9.9872
5.4877 | -34.1661 | -9.9974
5.4938 | -34.1664 | -10.0076
5.4999 | -34.1667 | -10.0179
5.506 | -34.1671 | -10.0283
5.5121 | -34.1674 | -10.0386
5.5183 | -34.1677 | -10.0487
5.5244 | -34.168 | -10.0594
5.5305 | -34.1684 | -10.0697
5.5366 | -34.1687 | -10.0798
5.5427 | -34.169 | -10.0904
5.5488 | -34.1693 | -10.1008
5.5549 | -34.1696 | -10.1112
5.561 | -34.17 | -10.1219
5.5671 | -34.1703 | -10.1321
5.5732 | -34.1707 | -10.1428
5.5793 | -34.171 | -10.1535
5.5854 | -34.1714 | -10.1639
5.5915 | -34.1717 | -10.1746
5.5976 | -34.172 | -10.1851
5.6037 | -34.1723 | -10.1956
5.6098 | -34.1726 | -10.206
5.6159 | -34.173 | -10.2162
5.622 | -34.1733 | -10.2267
5.6281 | -34.1737 | -10.2371
5.6342 | -34.174 | -10.2479
5.6403 | -34.1743 | -10.2584
5.6464 | -34.1746 | -10.2689
5.6525 | -34.1749 | -10.2795
5.6586 | -34.1753 | -10.2896
5.6648 | -34.1757 | -10.3001
5.6709 | -34.176 | -10.3107
5.677 | -34.1763 | -10.3212
5.6831 | -34.1767 | -10.3317
5.6892 | -34.177 | -10.3423
5.6953 | -34.1773 | -10.3527
5.7014 | -34.1776 | -10.3632
5.7075 | -34.1779 | -10.3737
5.7136 | -34.1782 | -10.384
5.7197 | -34.1785 | -10.3944
5.7258 | -34.1788 | -10.4046
5.7319 | -34.1792 | -10.4151
5.738 | -34.1795 | -10.4255
5.7441 | -34.1799 | -10.4359
5.7502 | -34.1802 | -10.4461
5.7563 | -34.1806 | -10.4565
5.7624 | -34.1809 | -10.4671
5.7685 | -34.1813 | -10.4776
5.7746 | -34.1816 | -10.4883
5.7807 | -34.1818 | -10.4988
5.7868 | -34.1821 | -10.5094
5.7929 | -34.1824 | -10.5198
5.799 | -34.1828 | -10.53
5.8052 | -34.1831 | -10.5403
5.8113 | -34.1834 | -10.5506
5.8174 | -34.1837 | -10.5608
5.8235 | -34.1841 | -10.5712
5.8296 | -34.1844 | -10.5818
5.8357 | -34.1846 | -10.592
5.8418 | -34.185 | -10.6026
5.8479 | -34.1853 | -10.6131
5.854 | -34.1856 | -10.6237
5.8601 | -34.1859 | -10.634
5.8662 | -34.1863 | -10.6446
5.8723 | -34.1866 | -10.6548
5.8784 | -34.1869 | -10.6651
5.8845 | -34.1872 | -10.6754
5.8906 | -34.1875 | -10.6858
5.8967 | -34.1878 | -10.696
5.9028 | -34.1882 | -10.7063
5.9089 | -34.1886 | -10.7165
5.915 | -34.1889 | -10.7267
5.9211 | -34.1892 | -10.737
5.9272 | -34.1895 | -10.7471
5.9333 | -34.1898 | -10.7573
5.9394 | -34.1902 | -10.7676
5.9455 | -34.1905 | -10.7778
5.9517 | -34.1909 | -10.788
5.9578 | -34.1913 | -10.7985
5.9639 | -34.1916 | -10.8085
5.97 | -34.192 | -10.8189
5.9761 | -34.1923 | -10.8291
5.9822 | -34.1926 | -10.8396
5.9883 | -34.193 | -10.85
5.9944 | -34.1933 | -10.8605
6.0005 | -34.1937 | -10.871
6.0066 | -34.194 | -10.8811
6.0127 | -34.1944 | -10.8911
6.0188 | -34.1947 | -10.9017
6.0249 | -34.195 | -10.9121
6.031 | -34.1953 | -10.9222
6.0371 | -34.1957 | -10.9326
6.0432 | -34.196 | -10.9428
6.0493 | -34.1963 | -10.953
6.0554 | -34.1967 | -10.9631
6.0615 | -34.197 | -10.9735
6.0676 | -34.1974 | -10.984
6.0737 | -34.1977 | -10.9944
6.0798 | -34.198 | -11.0047
6.0859 | -34.1984 | -11.015
6.0921 | -34.1987 | -11.025
6.0982 | -34.1991 | -11.0353
6.1043 | -34.1994 | -11.0455
6.1104 | -34.1998 | -11.056
6.1165 | -34.2002 | -11.0659
6.1226 | -34.2005 | -11.0761
6.1287 | -34.2008 | -11.0863
6.1348 | -34.2012 | -11.0966
6.1409 | -34.2015 | -11.1068
6.147 | -34.2019 | -11.117
6.1531 | -34.2022 | -11.1272
6.1592 | -34.2025 | -11.1374
6.1653 | -34.2028 | -11.1475
6.1714 | -34.2032 | -11.1578
6.1775 | -34.2035 | -11.1682
6.1836 | -34.2039 | -11.1785
6.1897 | -34.2042 | -11.1888
6.1958 | -34.2045 | -11.1993
6.2019 | -34.2049 | -11.2096
6.208 | -34.2053 | -11.2199
6.2141 | -34.2056 | -11.2301
6.2202 | -34.206 | -11.2405
6.2263 | -34.2064 | -11.2507
6.2325 | -34.2067 | -11.2607
6.2386 | -34.207 | -11.2706
6.2447 | -34.2074 | -11.2811
6.2508 | -34.2077 | -11.2914
6.2569 | -34.2081 | -11.3016
6.263 | -34.2084 | -11.3118
6.2691 | -34.2087 | -11.3219
6.2752 | -34.2091 | -11.3322
6.2813 | -34.2094 | -11.3425
6.2874 | -34.2097 | -11.3526
6.2935 | -34.2101 | -11.3628
6.2996 | -34.2104 | -11.3729
6.3057 | -34.2107 | -11.3832
6.3118 | -34.211 | -11.3936
6.3179 | -34.2114 | -11.4038
6.324 | -34.2117 | -11.4141
6.3301 | -34.2121 | -11.4243
6.3362 | -34.2124 | -11.4344
6.3423 | -34.2128 | -11.4448
6.3484 | -34.2131 | -11.4552
6.3545 | -34.2135 | -11.4651
6.3606 | -34.2138 | -11.475
6.3667 | -34.2141 | -11.4851
6.3728 | -34.2145 | -11.4954
6.379 | -34.2148 | -11.5056
6.3851 | -34.2152 | -11.5155
6.3912 | -34.2154 | -11.5257
6.3973 | -34.2158 | -11.536
6.4034 | -34.2161 | -11.546
6.4095 | -34.2164 | -11.5562
6.4156 | -34.2168 | -11.5664
6.4217 | -34.2171 | -11.5765
6.4278 | -34.2174 | -11.5868
6.4339 | -34.2178 | -11.5969
6.44 | -34.2182 | -11.6069
6.4461 | -34.2185 | -11.6173
6.4522 | -34.2188 | -11.6273
6.4583 | -34.2191 | -11.6376
6.4644 | -34.2194 | -11.6479
6.4705 | -34.2198 | -11.6577
6.4766 | -34.2201 | -11.668
6.4827 | -34.2204 | -11.6782
6.4888 | -34.2207 | -11.6882
6.4949 | -34.2211 | -11.698
6.501 | -34.2214 | -11.7083
6.5071 | -34.2218 | -11.7183
6.5132 | -34.2221 | -11.7285
6.5194 | -34.2225 | -11.7387
6.5255 | -34.2228 | -11.7486
6.5316 | -34.2231 | -11.7589
6.5377 | -34.2234 | -11.7691
6.5438 | -34.2237 | -11.7791
6.5499 | -34.2241 | -11.7891
6.556 | -34.2244 | -11.7992
6.5621 | -34.2248 | -11.8096
6.5682 | -34.2251 | -11.8198
6.5743 | -34.2255 | -11.8299
6.5804 | -34.2258 | -11.8402
6.5865 | -34.2261 | -11.8504
6.5926 | -34.2265 | -11.8604
6.5987 | -34.2268 | -11.8703
6.6048 | -34.2272 | -11.8805
6.6109 | -34.2275 | -11.8907
6.617 | -34.2278 | -11.901
6.6231 | -34.2281 | -11.9111
6.6292 | -34.2285 | -11.9213
6.6353 | -34.2288 | -11.9318
6.6414 | -34.2292 | -11.9418
6.6475 | -34.2295 | -11.952
6.6536 | -34.2298 | -11.9621
6.6597 | -34.2302 | -11.9719
6.6659 | -34.2305 | -11.9821
6.672 | -34.2309 | -11.9921
6.6781 | -34.2312 | -12.0023
6.6842 | -34.2315 | -12.0122
6.6903 | -34.2319 | -12.0223
6.6964 | -34.2323 | -12.0324
6.7025 | -34.2326 | -12.0423
6.7086 | -34.2329 | -12.0524
6.7147 | -34.2332 | -12.0626
6.7208 | -34.2336 | -12.0728
6.7269 | -34.2339 | -12.0829
6.733 | -34.2342 | -12.0928
6.7391 | -34.2346 | -12.1029
6.7452 | -34.2349 | -12.1128
6.7513 | -34.2352 | -12.1226
6.7574 | -34.2356 | -12.1328
6.7635 | -34.2359 | -12.1427
6.7696 | -34.2363 | -12.1527
6.7757 | -34.2367 | -12.1628
6.7818 | -34.237 | -12.1729
6.7879 | -34.2374 | -12.1828
6.794 | -34.2377 | -12.1927
6.8001 | -34.238 | -12.2028
6.8063 | -34.2383 | -12.2129
6.8124 | -34.2387 | -12.2231
6.8185 | -34.239 | -12.2334
6.8246 | -34.2393 | -12.2434
6.8307 | -34.2396 | -12.2533
6.8368 | -34.24 | -12.2636
6.8429 | -34.2403 | -12.2738
6.849 | -34.2407 | -12.2839
6.8551 | -34.241 | -12.2938
6.8612 | -34.2414 | -12.3039
6.8673 | -34.2417 | -12.3139
6.8734 | -34.242 | -12.3242
6.8795 | -34.2424 | -12.3342
6.8856 | -34.2427 | -12.3443
6.8917 | -34.2431 | -12.3543
6.8978 | -34.2434 | -12.3641
6.9039 | -34.2437 | -12.3739
6.91 | -34.2441 | -12.384
6.9161 | -34.2445 | -12.3939
6.9222 | -34.2448 | -12.4039
6.9283 | -34.2452 | -12.4137
6.9344 | -34.2455 | -12.4237
6.9405 | -34.2459 | -12.4336
6.9466 | -34.2462 | -12.4439
6.9528 | -34.2466 | -12.4541
6.9589 | -34.247 | -12.4643
6.965 | -34.2473 | -12.4745
6.9711 | -34.2477 | -12.4846
6.9772 | -34.2481 | -12.4946
6.9833 | -34.2484 | -12.5046
6.9894 | -34.2488 | -12.5149
6.9955 | -34.2491 | -12.5247
7.0016 | -34.2495 | -12.5347
7.0077 | -34.2498 | -12.5447
7.0138 | -34.2502 | -12.5547
7.0199 | -34.2505 | -12.5646
7.026 | -34.2509 | -12.5746
7.0321 | -34.2512 | -12.5844
7.0382 | -34.2516 | -12.5941
7.0443 | -34.252 | -12.6043
7.0504 | -34.2523 | -12.6143
7.0565 | -34.2526 | -12.6243
7.0626 | -34.253 | -12.6341
7.0687 | -34.2534 | -12.6437
7.0748 | -34.2537 | -12.6534
7.0809 | -34.2541 | -12.6633
7.087 | -34.2544 | -12.6735
7.0932 | -34.2548 | -12.6835
7.0993 | -34.2551 | -12.6934
7.1054 | -34.2555 | -12.7033
7.1115 | -34.2559 | -12.7131
7.1176 | -34.2562 | -12.7229
7.1237 | -34.2565 | -12.7328
7.1298 | -34.2569 | -12.7427
7.1359 | -34.2572 | -12.7527
7.142 | -34.2576 | -12.7623
7.1481 | -34.2579 | -12.7722
7.1542 | -34.2583 | -12.7821
7.1603 | -34.2586 | -12.7921
7.1664 | -34.259 | -12.802
7.1725 | -34.2594 | -12.8119
7.1786 | -34.2598 | -12.8218
7.1847 | -34.2601 | -12.8318
7.1908 | -34.2604 | -12.8419
7.1969 | -34.2608 | -12.852
7.203 | -34.2611 | -12.8617
7.2091 | -34.2615 | -12.8715
7.2152 | -34.2619 | -12.8814
7.2213 | -34.2622 | -12.8911
7.2274 | -34.2626 | -12.901
7.2335 | -34.2629 | -12.9108
7.2397 | -34.2633 | -12.9206
7.2458 | -34.2637 | -12.9302
7.2519 | -34.2641 | -12.9402
7.258 | -34.2644 | -12.9503
7.2641 | -34.2648 | -12.9602
7.2702 | -34.2651 | -12.9702
7.2763 | -34.2655 | -12.9802
7.2824 | -34.2659 | -12.9901
7.2885 | -34.2662 | -13
7.2946 | -34.2666 | -13.0099
7.3007 | -34.2669 | -13.0199
7.3068 | -34.2673 | -13.03
7.3129 | -34.2676 | -13.04
7.319 | -34.268 | -13.0498
7.3251 | -34.2684 | -13.0595
7.3312 | -34.2687 | -13.0694
7.3373 | -34.2691 | -13.0791
7.3434 | -34.2695 | -13.0889
7.3495 | -34.2699 | -13.0986
7.3556 | -34.2702 | -13.1085
7.3617 | -34.2706 | -13.1186
7.3678 | -34.271 | -13.1284
7.3739 | -34.2713 | -13.1385
7.3801 | -34.2717 | -13.1486
7.3862 | -34.2721 | -13.1584
7.3923 | -34.2724 | -13.1685
7.3984 | -34.2728 | -13.1781
7.4045 | -34.2732 | -13.1879
7.4106 | -34.2736 | -13.1977
7.4167 | -34.2739 | -13.2077
7.4228 | -34.2743 | -13.2177
7.4289 | -34.2747 | -13.2275
7.435 | -34.275 | -13.2371
7.4411 | -34.2754 | -13.247
7.4472 | -34.2758 | -13.257
7.4533 | -34.2761 | -13.2665
7.4594 | -34.2765 | -13.2766
7.4655 | -34.2769 | -13.2863
7.4716 | -34.2773 | -13.2964
7.4777 | -34.2776 | -13.3063
7.4838 | -34.278 | -13.3159
7.4899 | -34.2783 | -13.3256
7.496 | -34.2787 | -13.3352
7.5021 | -34.2791 | -13.345
7.5082 | -34.2794 | -13.3549
7.5143 | -34.2798 | -13.3648
7.5204 | -34.2802 | -13.3748
7.5266 | -34.2805 | -13.3844
7.5327 | -34.2809 | -13.3942
7.5388 | -34.2812 | -13.4043
7.5449 | -34.2816 | -13.4144
7.551 | -34.282 | -13.4241
7.5571 | -34.2824 | -13.434
7.5632 | -34.2828 | -13.4439
7.5693 | -34.2832 | -13.4536
7.5754 | -34.2835 | -13.4633
7.5815 | -34.2839 | -13.4732
7.5876 | -34.2842 | -13.4832
7.5937 | -34.2846 | -13.4932
7.5998 | -34.285 | -13.503
7.6059 | -34.2854 | -13.5131
7.612 | -34.2858 | -13.5229
7.6181 | -34.2862 | -13.5326
7.6242 | -34.2865 | -13.5428
7.6303 | -34.2869 | -13.5528
7.6364 | -34.2872 | -13.5628
7.6425 | -34.2876 | -13.5727
7.6486 | -34.288 | -13.5826
7.6547 | -34.2884 | -13.5923
7.6608 | -34.2887 | -13.6021
7.667 | -34.2891 | -13.612
7.6731 | -34.2895 | -13.622
7.6792 | -34.2898 | -13.6321
7.6853 | -34.2902 | -13.642
7.6914 | -34.2906 | -13.652
7.6975 | -34.291 | -13.662
7.7036 | -34.2913 | -13.6719
7.7097 | -34.2917 | -13.6819
7.7158 | -34.2921 | -13.6918
7.7219 | -34.2925 | -13.7015
7.728 | -34.2929 | -13.7115
7.7341 | -34.2932 | -13.7214
7.7402 | -34.2936 | -13.7314
7.7463 | -34.294 | -13.7413
7.7524 | -34.2944 | -13.7514
7.7585 | -34.2947 | -13.7613
7.7646 | -34.2951 | -13.7712
7.7707 | -34.2954 | -13.7812
7.7768 | -34.2958 | -13.7913
7.7829 | -34.2962 | -13.8015
7.789 | -34.2965 | -13.8116
7.7951 | -34.2969 | -13.8215
7.8012 | -34.2973 | -13.8314
7.8073 | -34.2977 | -13.8415
7.8135 | -34.2981 | -13.8514
7.8196 | -34.2985 | -13.8615
7.8257 | -34.2988 | -13.8716
7.8318 | -34.2992 | -13.8816
7.8379 | -34.2996 | -13.8914
7.844 | -34.2999 | -13.9013
7.8501 | -34.3003 | -13.9114
7.8562 | -34.3007 | -13.9216
7.8623 | -34.3011 | -13.9314
7.8684 | -34.3014 | -13.9414
7.8745 | -34.3018 | -13.9515
7.8806 | -34.3021 | -13.9615
7.8867 | -34.3025 | -13.9713
7.8928 | -34.3029 | -13.981
7.8989 | -34.3033 | -13.9909
7.905 | -34.3037 | -14.001
7.9111 | -34.304 | -14.011
7.9172 | -34.3044 | -14.0209
7.9233 | -34.3047 | -14.0309
7.9294 | -34.3051 | -14.041
7.9355 | -34.3055 | -14.0507
7.9416 | -34.3058 | -14.0607
7.9477 | -34.3062 | -14.0705
7.9539 | -34.3065 | -14.0803
7.96 | -34.3069 | -14.0903
7.9661 | -34.3073 | -14.1003
7.9722 | -34.3077 | -14.1101
7.9783 | -34.3081 | -14.1203
7.9844 | -34.3085 | -14.1301
7.9905 | -34.3088 | -14.1403
7.9966 | -34.3092 | -14.1503
8.0027 | -34.3095 | -14.1603
8.0088 | -34.3099 | -14.1702
8.0149 | -34.3103 | -14.1802
8.021 | -34.3106 | -14.1902
8.0271 | -34.311 | -14.2004
8.0332 | -34.3114 | -14.2102
8.0393 | -34.3118 | -14.2203
8.0454 | -34.3121 | -14.2305
8.0515 | -34.3126 | -14.2405
8.0576 | -34.3129 | -14.2507
8.0637 | -34.3133 | -14.2605
8.0698 | -34.3136 | -14.2703
8.0759 | -34.314 | -14.2802
8.082 | -34.3144 | -14.2901
8.0881 | -34.3147 | -14.3001
8.0942 | -34.3151 | -14.3101
8.1004 | -34.3155 | -14.3202
8.1065 | -34.3159 | -14.33
8.1126 | -34.3162 | -14.3401
8.1187 | -34.3166 | -14.3501
8.1248 | -34.317 | -14.3599
8.1309 | -34.3173 | -14.3697
8.137 | -34.3177 | -14.3794
8.1431 | -34.3181 | -14.3896
8.1492 | -34.3185 | -14.3995
8.1553 | -34.3188 | -14.4096
8.1614 | -34.3192 | -14.4197
8.1675 | -34.3195 | -14.4297
8.1736 | -34.3199 | -14.4398
8.1797 | -34.3203 | -14.4498
8.1858 | -34.3206 | -14.4599
8.1919 | -34.321 | -14.4698
8.198 | -34.3214 | -14.4796
8.2041 | -34.3218 | -14.4897
8.2102 | -34.3221 | -14.4996
8.2163 | -34.3225 | -14.5094
8.2224 | -34.3229 | -14.5194
8.2285 | -34.3232 | -14.5291
8.2346 | -34.3236 | -14.539
8.2408 | -34.324 | -14.5488
8.2469 | -34.3244 | -14.5588
8.253 | -34.3248 | -14.5687
8.2591 | -34.3251 | -14.5787
8.2652 | -34.3255 | -14.5887
8.2713 | -34.3259 | -14.5988
8.2774 | -34.3263 | -14.6088
8.2835 | -34.3266 | -14.619
8.2896 | -34.327 | -14.6287
8.2957 | -34.3274 | -14.6388
8.3018 | -34.3278 | -14.6488
8.3079 | -34.3282 | -14.6587
8.314 | -34.3285 | -14.6686
8.3201 | -34.3289 | -14.6785
8.3262 | -34.3293 | -14.6885
8.3323 | -34.3297 | -14.6984
8.3384 | -34.33 | -14.7085
8.3445 | -34.3304 | -14.7183
8.3506 | -34.3308 | -14.7283
8.3567 | -34.3312 | -14.7383
8.3628 | -34.3315 | -14.7483
8.3689 | -34.3319 | -14.7581
8.375 | -34.3323 | -14.7681
8.3812 | -34.3326 | -14.7779
8.3873 | -34.333 | -14.7878
8.3934 | -34.3334 | -14.7979
8.3995 | -34.3337 | -14.8077
8.4056 | -34.3341 | -14.8176
8.4117 | -34.3345 | -14.8276
8.4178 | -34.3348 | -14.8375
8.4239 | -34.3352 | -14.8473
8.43 | -34.3356 | -14.8573
8.4361 | -34.3359 | -14.8673
8.4422 | -34.3363 | -14.8771
8.4483 | -34.3367 | -14.887
8.4544 | -34.3371 | -14.8969
8.4605 | -34.3374 | -14.9067
8.4666 | -34.3378 | -14.9168
8.4727 | -34.3382 | -14.9268
8.4788 | -34.3385 | -14.9367
8.4849 | -34.3389 | -14.9466
8.491 | -34.3393 | -14.9565
8.4971 | -34.3397 | -14.9663
8.5032 | -34.34 | -14.9761
8.5093 | -34.3404 | -14.9858
8.5154 | -34.3408 | -14.9958
8.5215 | -34.3412 | -15.0056
8.5277 | -34.3415 | -15.0155
8.5338 | -34.3419 | -15.0254
8.5399 | -34.3423 | -15.0353
8.546 | -34.3427 | -15.045
8.5521 | -34.3431 | -15.0549
8.5582 | -34.3434 | -15.0648
8.5643 | -34.3438 | -15.0744
8.5704 | -34.3442 | -15.0841
8.5765 | -34.3445 | -15.0941
8.5826 | -34.3449 | -15.1036
8.5887 | -34.3453 | -15.1134
8.5948 | -34.3457 | -15.1234
8.6009 | -34.3461 | -15.133
8.607 | -34.3465 | -15.143
8.6131 | -34.3469 | -15.1529
8.6192 | -34.3473 | -15.1626
8.6253 | -34.3477 | -15.1724
8.6314 | -34.348 | -15.1822
8.6375 | -34.3484 | -15.1921
8.6436 | -34.3488 | -15.2019
8.6497 | -34.3492 | -15.2117
8.6558 | -34.3495 | -15.2215
8.6619 | -34.3499 | -15.2314
8.6681 | -34.3503 | -15.2414
8.6742 | -34.3507 | -15.2512
8.6803 | -34.351 | -15.261
8.6864 | -34.3514 | -15.2705
8.6925 | -34.3518 | -15.2801
8.6986 | -34.3522 | -15.29
8.7047 | -34.3526 | -15.3
8.7108 | -34.3529 | -15.3099
8.7169 | -34.3533 | -15.3197
8.723 | -34.3537 | -15.3294
8.7291 | -34.3541 | -15.339
8.7352 | -34.3545 | -15.349
8.7413 | -34.3548 | -15.3589
8.7474 | -34.3552 | -15.3686
8.7535 | -34.3556 | -15.3784
8.7596 | -34.356 | -15.3883
8.7657 | -34.3564 | -15.3982
8.7718 | -34.3567 | -15.4081
8.7779 | -34.3571 | -15.418
8.784 | -34.3575 | -15.4279
8.7901 | -34.3579 | -15.4377
8.7962 | -34.3583 | -15.4473
8.8023 | -34.3587 | -15.457
8.8084 | -34.3591 | -15.4668
8.8146 | -34.3595 | -15.4766
8.8207 | -34.3598 | -15.4863
8.8268 | -34.3602 | -15.4958
8.8329 | -34.3606 | -15.5056
8.839 | -34.361 | -15.5155
8.8451 | -34.3614 | -15.5251
8.8512 | -34.3618 | -15.5348
8.8573 | -34.3621 | -15.5446
8.8634 | -34.3625 | -15.5542
8.8695 | -34.3629 | -15.5642
8.8756 | -34.3633 | -15.574
8.8817 | -34.3637 | -15.5836
8.8878 | -34.3641 | -15.5936
8.8939 | -34.3645 | -15.6034
8.9 | -34.3649 | -15.6131
8.9061 | -34.3652 | -15.6228
8.9122 | -34.3656 | -15.6326
8.9183 | -34.366 | -15.6424
8.9244 | -34.3664 | -15.6523
8.9305 | -34.3669 | -15.6619
8.9366 | -34.3673 | -15.6716
8.9427 | -34.3677 | -15.6815
8.9488 | -34.3681 | -15.6912
8.955 | -34.3684 | -15.7009
8.9611 | -34.3688 | -15.7108
8.9672 | -34.3693 | -15.7206
8.9733 | -34.3697 | -15.7301
8.9794 | -34.3701 | -15.7398
8.9855 | -34.3705 | -15.7496
8.9916 | -34.3709 | -15.7595
8.9977 | -34.3713 | -15.7692
9.0038 | -34.3717 | -15.7789
9.0099 | -34.3721 | -15.7886
9.016 | -34.3725 | -15.7984
9.0221 | -34.3729 | -15.8079
9.0282 | -34.3733 | -15.8175
9.0343 | -34.3737 | -15.8272
9.0404 | -34.3741 | -15.8369
9.0465 | -34.3745 | -15.8468
9.0526 | -34.375 | -15.8565
9.0587 | -34.3754 | -15.8662
9.0648 | -34.3758 | -15.8758
9.0709 | -34.3762 | -15.8855
9.077 | -34.3766 | -15.8953
9.0831 | -34.377 | -15.9051
9.0892 | -34.3774 | -15.9149
9.0953 | -34.3778 | -15.9248
9.1015 | -34.3782 | -15.9345
9.1076 | -34.3787 | -15.944
9.1137 | -34.3791 | -15.9537
9.1198 | -34.3795 | -15.9634
9.1259 | -34.3799 | -15.9731
9.132 | -34.3803 | -15.983
9.1381 | -34.3808 | -15.9926
9.1442 | -34.3812 | -16.002
9.1503 | -34.3816 | -16.0118
9.1564 | -34.382 | -16.0214
9.1625 | -34.3824 | -16.0312
9.1686 | -34.3828 | -16.0409
9.1747 | -34.3832 | -16.0506
9.1808 | -34.3837 | -16.0603
9.1869 | -34.3841 | -16.07
9.193 | -34.3845 | -16.0797
9.1991 | -34.3849 | -16.0894
9.2052 | -34.3853 | -16.0991
9.2113 | -34.3858 | -16.1089
9.2174 | -34.3862 | -16.1186
9.2235 | -34.3866 | -16.1283
9.2296 | -34.3871 | -16.138
9.2357 | -34.3875 | -16.1478
9.2419 | -34.3879 | -16.1575
9.248 | -34.3884 | -16.1672
9.2541 | -34.3888 | -16.1769
9.2602 | -34.3892 | -16.1868
9.2663 | -34.3896 | -16.1962
9.2724 | -34.3901 | -16.2059
9.2785 | -34.3905 | -16.2155
9.2846 | -34.3909 | -16.2251
9.2907 | -34.3913 | -16.2348
9.2968 | -34.3918 | -16.2446
9.3029 | -34.3922 | -16.2544
9.309 | -34.3926 | -16.2641
9.3151 | -34.393 | -16.2736
9.3212 | -34.3935 | -16.2832
9.3273 | -34.3939 | -16.2928
9.3334 | -34.3943 | -16.3024
9.3395 | -34.3948 | -16.3119
9.3456 | -34.3952 | -16.3214
9.3517 | -34.3956 | -16.3312
9.3578 | -34.396 | -16.3409
9.3639 | -34.3965 | -16.3504
9.37 | -34.3969 | -16.36
9.3761 | -34.3973 | -16.3699
9.3822 | -34.3978 | -16.3796
9.3884 | -34.3982 | -16.389
9.3945 | -34.3986 | -16.3988
9.4006 | -34.3991 | -16.4084
9.4067 | -34.3995 | -16.4181
9.4128 | -34.3999 | -16.4279
9.4189 | -34.4004 | -16.4377
9.425 | -34.4008 | -16.4473
9.4311 | -34.4013 | -16.4569
9.4372 | -34.4017 | -16.4665
9.4433 | -34.4022 | -16.4762
9.4494 | -34.4026 | -16.4857
9.4555 | -34.4031 | -16.4955
9.4616 | -34.4035 | -16.5051
9.4677 | -34.4039 | -16.5147
9.4738 | -34.4044 | -16.5244
9.4799 | -34.4048 | -16.5341
9.486 | -34.4052 | -16.5437
9.4921 | -34.4057 | -16.5533
9.4982 | -34.4061 | -16.5628
9.5043 | -34.4066 | -16.5724
9.5104 | -34.407 | -16.5819
9.5165 | -34.4074 | -16.5916
9.5226 | -34.4078 | -16.6014
9.5288 | -34.4083 | -16.6108
9.5349 | -34.4087 | -16.6206
9.541 | -34.4092 | -16.63
9.5471 | -34.4096 | -16.6396
9.5532 | -34.41 | -16.6493
9.5593 | -34.4105 | -16.6591
9.5654 | -34.4109 | -16.6686
9.5715 | -34.4113 | -16.6782
9.5776 | -34.4118 | -16.688
9.5837 | -34.4122 | -16.6975
9.5898 | -34.4126 | -16.7072
9.5959 | -34.413 | -16.7169
9.602 | -34.4135 | -16.7264
9.6081 | -34.4139 | -16.736
9.6142 | -34.4144 | -16.7455
9.6203 | -34.4148 | -16.7552
9.6264 | -34.4152 | -16.7647
9.6325 | -34.4156 | -16.7744
9.6386 | -34.4161 | -16.7839
9.6447 | -34.4165 | -16.7936
9.6508 | -34.4169 | -16.8031
9.6569 | -34.4174 | -16.8127
9.663 | -34.4178 | -16.8224
9.6691 | -34.4183 | -16.8321
9.6753 | -34.4188 | -16.8416
9.6814 | -34.4192 | -16.851
9.6875 | -34.4196 | -16.8604
9.6936 | -34.4201 | -16.8698
9.6997 | -34.4205 | -16.8796
9.7058 | -34.421 | -16.8891
9.7119 | -34.4214 | -16.8988
9.718 | -34.4219 | -16.9085
9.7241 | -34.4223 | -16.918
9.7302 | -34.4227 | -16.9276
9.7363 | -34.4232 | -16.9371
9.7424 | -34.4236 | -16.9466
9.7485 | -34.4241 | -16.9562
9.7546 | -34.4245 | -16.9656
9.7607 | -34.4249 | -16.9754
9.7668 | -34.4254 | -16.9851
9.7729 | -34.4258 | -16.9946
9.779 | -34.4263 | -17.0039
9.7851 | -34.4267 | -17.0137
9.7912 | -34.4272 | -17.0232
9.7973 | -34.4277 | -17.0329
9.8034 | -34.4281 | -17.0422
9.8095 | -34.4286 | -17.052
9.8157 | -34.429 | -17.0616
9.8218 | -34.4295 | -17.0713
9.8279 | -34.4299 | -17.0807
9.834 | -34.4303 | -17.0904
9.8401 | -34.4308 | -17.1
9.8462 | -34.4312 | -17.1093
9.8523 | -34.4317 | -17.1189
9.8584 | -34.4321 | -17.1284
9.8645 | -34.4325 | -17.1382
9.8706 | -34.4329 | -17.1478
9.8767 | -34.4334 | -17.1573
9.8828 | -34.4338 | -17.1667
9.8889 | -34.4343 | -17.1761
9.895 | -34.4347 | -17.1857
9.9011 | -34.4351 | -17.1953
9.9072 | -34.4356 | -17.2048
9.9133 | -34.436 | -17.2143
9.9194 | -34.4364 | -17.2238
9.9255 | -34.4369 | -17.2332
9.9316 | -34.4373 | -17.2428
9.9377 | -34.4378 | -17.2523
9.9438 | -34.4382 | -17.2619
9.9499 | -34.4386 | -17.2716
9.956 | -34.4391 | -17.2812
9.9622 | -34.4395 | -17.2908
9.9683 | -34.4399 | -17.3004
9.9744 | -34.4404 | -17.3101
9.9805 | -34.4408 | -17.3196
9.9866 | -34.4413 | -17.3291
9.9927 | -34.4417 | -17.3385
9.9988 | -34.4421 | -17.3482
10.0049 | -34.4426 | -17.3577
10.011 | -34.443 | -17.3673
10.0171 | -34.4435 | -17.3767
10.0232 | -34.4439 | -17.3861
10.0293 | -34.4444 | -17.3957
10.0354 | -34.4448 | -17.4053
10.0415 | -34.4453 | -17.4148
10.0476 | -34.4457 | -17.4245
10.0537 | -34.4461 | -17.4341
10.0598 | -34.4466 | -17.4436
10.0659 | -34.447 | -17.4532
10.072 | -34.4475 | -17.4625
10.0781 | -34.4479 | -17.472
10.0842 | -34.4483 | -17.4815
10.0903 | -34.4488 | -17.491
10.0964 | -34.4492 | -17.5002
10.1026 | -34.4496 | -17.5098
10.1087 | -34.4501 | -17.5193
10.1148 | -34.4505 | -17.5289
10.1209 | -34.4509 | -17.5383
10.127 | -34.4514 | -17.5476
10.1331 | -34.4518 | -17.557
10.1392 | -34.4522 | -17.5663
10.1453 | -34.4527 | -17.5758
10.1514 | -34.4531 | -17.5853
10.1575 | -34.4536 | -17.5949
10.1636 | -34.454 | -17.6042
10.1697 | -34.4544 | -17.6138
10.1758 | -34.4549 | -17.6231
10.1819 | -34.4553 | -17.6324
10.188 | -34.4557 | -17.642
10.1941 | -34.4562 | -17.6515
10.2002 | -34.4566 | -17.6609
10.2063 | -34.457 | -17.6704
10.2124 | -34.4575 | -17.6799
10.2185 | -34.4579 | -17.6895
10.2246 | -34.4584 | -17.6988
10.2307 | -34.4588 | -17.7084
10.2368 | -34.4592 | -17.7175
10.2429 | -34.4596 | -17.727
10.2491 | -34.4601 | -17.7365
10.2552 | -34.4605 | -17.746
10.2613 | -34.461 | -17.7556
10.2674 | -34.4614 | -17.7651
10.2735 | -34.4618 | -17.7747
10.2796 | -34.4623 | -17.7842
10.2857 | -34.4627 | -17.7938
10.2918 | -34.4632 | -17.8034
10.2979 | -34.4636 | -17.8128
10.304 | -34.464 | -17.8221
10.3101 | -34.4645 | -17.8316
10.3162 | -34.4649 | -17.8412
10.3223 | -34.4653 | -17.8505
10.3284 | -34.4658 | -17.8599
10.3345 | -34.4662 | -17.8693
10.3406 | -34.4666 | -17.8787
10.3467 | -34.4671 | -17.888
10.3528 | -34.4675 | -17.8974
10.3589 | -34.4679 | -17.9068
10.365 | -34.4684 | -17.9164
10.3711 | -34.4688 | -17.9259
10.3772 | -34.4692 | -17.9354
10.3833 | -34.4697 | -17.9448
10.3895 | -34.4701 | -17.9543
10.3956 | -34.4706 | -17.9638
10.4017 | -34.471 | -17.973
10.4078 | -34.4714 | -17.9824
10.4139 | -34.4719 | -17.9919
10.42 | -34.4723 | -18.0015
10.4261 | -34.4728 | -18.0108
10.4322 | -34.4732 | -18.0202
10.4383 | -34.4736 | -18.0296
10.4444 | -34.474 | -18.0389
10.4505 | -34.4744 | -18.0484
10.4566 | -34.4749 | -18.0577
10.4627 | -34.4753 | -18.067
10.4688 | -34.4757 | -18.0763
10.4749 | -34.4762 | -18.0858
10.481 | -34.4766 | -18.0952
10.4871 | -34.4771 | -18.1045
10.4932 | -34.4775 | -18.1138
10.4993 | -34.4779 | -18.1233
10.5054 | -34.4784 | -18.1327
10.5115 | -34.4788 | -18.1421
10.5176 | -34.4793 | -18.1514
10.5237 | -34.4797 | -18.1609
10.5298 | -34.4801 | -18.1702
10.536 | -34.4806 | -18.1797
10.5421 | -34.481 | -18.1889
10.5482 | -34.4814 | -18.1984
10.5543 | -34.4819 | -18.2076
10.5604 | -34.4823 | -18.2168
10.5665 | -34.4828 | -18.2261
10.5726 | -34.4832 | -18.2355
10.5787 | -34.4836 | -18.2448
10.5848 | -34.4841 | -18.2538
10.5909 | -34.4845 | -18.2633
10.597 | -34.4849 | -18.2728
10.6031 | -34.4854 | -18.2823
10.6092 | -34.4858 | -18.2916
10.6153 | -34.4862 | -18.3011
10.6214 | -34.4867 | -18.3103
10.6275 | -34.4871 | -18.3196
10.6336 | -34.4875 | -18.329
10.6397 | -34.4879 | -18.3384
10.6458 | -34.4884 | -18.3477
10.6519 | -34.4888 | -18.3572
10.658 | -34.4892 | -18.3664
10.6641 | -34.4897 | -18.3759
10.6702 | -34.4901 | -18.3853
10.6764 | -34.4905 | -18.3948
10.6825 | -34.491 | -18.404
10.6886 | -34.4914 | -18.4132
10.6947 | -34.4919 | -18.4225
10.7008 | -34.4923 | -18.4319
10.7069 | -34.4928 | -18.4413
10.713 | -34.4932 | -18.4507
10.7191 | -34.4936 | -18.46
10.7252 | -34.4941 | -18.469
10.7313 | -34.4945 | -18.4783
10.7374 | -34.4949 | -18.4877
10.7435 | -34.4954 | -18.4971
10.7496 | -34.4958 | -18.5062
10.7557 | -34.4962 | -18.5157
10.7618 | -34.4967 | -18.5247
10.7679 | -34.4971 | -18.534
10.774 | -34.4975 | -18.5435
10.7801 | -34.498 | -18.5528
10.7862 | -34.4984 | -18.5619
10.7923 | -34.4988 | -18.5714
10.7984 | -34.4993 | -18.5805
10.8045 | -34.4997 | -18.5897
10.8106 | -34.5001 | -18.5991
10.8168 | -34.5006 | -18.6083
10.8229 | -34.501 | -18.6176
10.829 | -34.5014 | -18.6268
10.8351 | -34.5019 | -18.6358
10.8412 | -34.5023 | -18.645
10.8473 | -34.5028 | -18.6541
10.8534 | -34.5032 | -18.6634
10.8595 | -34.5037 | -18.6727
10.8656 | -34.5041 | -18.6817
10.8717 | -34.5045 | -18.691
10.8778 | -34.505 | -18.7003
10.8839 | -34.5054 | -18.7095
10.89 | -34.5059 | -18.7188
10.8961 | -34.5063 | -18.728
10.9022 | -34.5067 | -18.7372
10.9083 | -34.5072 | -18.7463
10.9144 | -34.5076 | -18.7555
10.9205 | -34.508 | -18.7647
10.9266 | -34.5085 | -18.774
10.9327 | -34.5089 | -18.7832
10.9388 | -34.5094 | -18.7922
10.9449 | -34.5098 | -18.8015
10.951 | -34.5103 | -18.8107
10.9571 | -34.5107 | -18.82
10.9633 | -34.5111 | -18.8293
10.9694 | -34.5116 | -18.8387
10.9755 | -34.512 | -18.8479
10.9816 | -34.5124 | -18.8573
10.9877 | -34.5129 | -18.8664
10.9938 | -34.5133 | -18.8758
10.9999 | -34.5138 | -18.8849
11.006 | -34.5142 | -18.8939
11.0121 | -34.5147 | -18.9031
11.0182 | -34.5151 | -18.9123
11.0243 | -34.5156 | -18.9215
11.0304 | -34.516 | -18.9307
11.0365 | -34.5164 | -18.9398
11.0426 | -34.5169 | -18.9489
11.0487 | -34.5174 | -18.9583
11.0548 | -34.5178 | -18.9673
11.0609 | -34.5183 | -18.9766
11.067 | -34.5187 | -18.9858
11.0731 | -34.5192 | -18.9949
11.0792 | -34.5196 | -19.004
11.0853 | -34.5201 | -19.0132
11.0914 | -34.5205 | -19.0225
11.0975 | -34.521 | -19.0316
11.1037 | -34.5214 | -19.0408
11.1098 | -34.5218 | -19.0501
11.1159 | -34.5223 | -19.059
11.122 | -34.5227 | -19.0681
11.1281 | -34.5232 | -19.0773
11.1342 | -34.5236 | -19.0866
11.1403 | -34.5241 | -19.0959
11.1464 | -34.5245 | -19.1051
11.1525 | -34.5249 | -19.114
11.1586 | -34.5254 | -19.1233
11.1647 | -34.5258 | -19.1324
11.1708 | -34.5263 | -19.1416
11.1769 | -34.5267 | -19.1507
11.183 | -34.5271 | -19.1598
11.1891 | -34.5276 | -19.1689
11.1952 | -34.5281 | -19.1781
11.2013 | -34.5285 | -19.1873
11.2074 | -34.529 | -19.1964
11.2135 | -34.5294 | -19.2057
11.2196 | -34.5299 | -19.2147
11.2257 | -34.5303 | -19.2239
11.2318 | -34.5308 | -19.233
11.2379 | -34.5312 | -19.2423
11.244 | -34.5317 | -19.2515
11.2502 | -34.5322 | -19.2606
11.2563 | -34.5326 | -19.2698
11.2624 | -34.5331 | -19.2788
11.2685 | -34.5335 | -19.2878
11.2746 | -34.534 | -19.2969
11.2807 | -34.5344 | -19.3062
11.2868 | -34.5349 | -19.3154
11.2929 | -34.5353 | -19.3243
11.299 | -34.5358 | -19.3335
11.3051 | -34.5362 | -19.3429
11.3112 | -34.5367 | -19.3521
11.3173 | -34.5371 | -19.3611
11.3234 | -34.5376 | -19.3699
11.3295 | -34.5381 | -19.3792
11.3356 | -34.5385 | -19.3885
11.3417 | -34.539 | -19.3975
11.3478 | -34.5394 | -19.4068
11.3539 | -34.5398 | -19.416
11.36 | -34.5403 | -19.425
11.3661 | -34.5407 | -19.434
11.3722 | -34.5412 | -19.4429
11.3783 | -34.5416 | -19.4521
11.3844 | -34.5421 | -19.4612
11.3906 | -34.5426 | -19.4703
11.3967 | -34.543 | -19.4793
11.4028 | -34.5435 | -19.4885
11.4089 | -34.5439 | -19.4975
11.415 | -34.5443 | -19.5066
11.4211 | -34.5448 | -19.5158
11.4272 | -34.5452 | -19.5249
11.4333 | -34.5457 | -19.5338
11.4394 | -34.5461 | -19.5429
11.4455 | -34.5465 | -19.5519
11.4516 | -34.547 | -19.5611
11.4577 | -34.5474 | -19.5702
11.4638 | -34.5479 | -19.579
11.4699 | -34.5484 | -19.5882
11.476 | -34.5488 | -19.5974
11.4821 | -34.5493 | -19.6064
11.4882 | -34.5497 | -19.6156
11.4943 | -34.5502 | -19.6248
11.5004 | -34.5506 | -19.6337
11.5065 | -34.5511 | -19.6426
11.5126 | -34.5515 | -19.6516
11.5187 | -34.552 | -19.6607
11.5248 | -34.5524 | -19.67
11.5309 | -34.5529 | -19.679
11.5371 | -34.5534 | -19.6879
11.5432 | -34.5538 | -19.697
11.5493 | -34.5543 | -19.7063
11.5554 | -34.5547 | -19.7154
11.5615 | -34.5552 | -19.7246
11.5676 | -34.5556 | -19.7337
11.5737 | -34.5561 | -19.7426
11.5798 | -34.5565 | -19.7515
11.5859 | -34.557 | -19.7605
11.592 | -34.5574 | -19.7695
11.5981 | -34.5579 | -19.7784
11.6042 | -34.5584 | -19.7875
11.6103 | -34.5588 | -19.7964
11.6164 | -34.5593 | -19.8053
11.6225 | -34.5597 | -19.8143
11.6286 | -34.5602 | -19.8233
11.6347 | -34.5606 | -19.8324
11.6408 | -34.5611 | -19.8413
11.6469 | -34.5615 | -19.8503
11.653 | -34.562 | -19.8593
11.6591 | -34.5625 | -19.8683
11.6652 | -34.5629 | -19.8772
11.6713 | -34.5634 | -19.8862
11.6775 | -34.5638 | -19.895
11.6836 | -34.5643 | -19.904
11.6897 | -34.5648 | -19.913
11.6958 | -34.5652 | -19.9221
11.7019 | -34.5657 | -19.9313
11.708 | -34.5661 | -19.9403
11.7141 | -34.5666 | -19.9491
11.7202 | -34.5671 | -19.9581
11.7263 | -34.5675 | -19.9672
11.7324 | -34.5679 | -19.9763
11.7385 | -34.5684 | -19.9854
11.7446 | -34.5688 | -19.9943
11.7507 | -34.5693 | -20.0031
11.7568 | -34.5698 | -20.0121
11.7629 | -34.5702 | -20.0211
11.769 | -34.5707 | -20.03
11.7751 | -34.5711 | -20.039
11.7812 | -34.5716 | -20.0478
11.7873 | -34.572 | -20.0568
11.7934 | -34.5725 | -20.0658
11.7995 | -34.573 | -20.0747
11.8056 | -34.5734 | -20.0836
11.8117 | -34.5739 | -20.0925
11.8178 | -34.5744 | -20.1013
11.824 | -34.5748 | -20.1102
11.8301 | -34.5753 | -20.1191
11.8362 | -34.5757 | -20.1279
11.8423 | -34.5762 | -20.1369
11.8484 | -34.5766 | -20.1458
11.8545 | -34.5771 | -20.1546
11.8606 | -34.5776 | -20.1635
11.8667 | -34.578 | -20.1725
11.8728 | -34.5785 | -20.1813
11.8789 | -34.5789 | -20.1903
11.885 | -34.5794 | -20.1992
11.8911 | -34.5799 | -20.208
11.8972 | -34.5803 | -20.2169
11.9033 | -34.5808 | -20.2258
11.9094 | -34.5812 | -20.2347
11.9155 | -34.5817 | -20.2436
11.9216 | -34.5822 | -20.2523
11.9277 | -34.5827 | -20.2612
11.9338 | -34.5831 | -20.2701
11.9399 | -34.5836 | -20.2791
11.946 | -34.584 | -20.288
11.9521 | -34.5845 | -20.2968
11.9582 | -34.585 | -20.3056
11.9644 | -34.5854 | -20.3144
11.9705 | -34.5859 | -20.3233
11.9766 | -34.5863 | -20.3322
11.9827 | -34.5868 | -20.3411
11.9888 | -34.5873 | -20.3499
11.9949 | -34.5878 | -20.3585
12.001 | -34.5882 | -20.3675
12.0071 | -34.5887 | -20.3763
12.0132 | -34.5891 | -20.3851
12.0193 | -34.5896 | -20.3941
12.0254 | -34.59 | -20.4029
12.0315 | -34.5905 | -20.4115
12.0376 | -34.591 | -20.4203
12.0437 | -34.5914 | -20.4292
12.0498 | -34.5919 | -20.4381
12.0559 | -34.5924 | -20.4469
12.062 | -34.5929 | -20.4555
12.0681 | -34.5933 | -20.4644
12.0742 | -34.5938 | -20.4731
12.0803 | -34.5942 | -20.4821
12.0864 | -34.5947 | -20.491
12.0925 | -34.5952 | -20.4998
12.0986 | -34.5957 | -20.5085
12.1047 | -34.5961 | -20.5173
12.1109 | -34.5966 | -20.526
12.117 | -34.597 | -20.5348
12.1231 | -34.5975 | -20.5436
12.1292 | -34.598 | -20.5524
12.1353 | -34.5984 | -20.5611
12.1414 | -34.5989 | -20.57
12.1475 | -34.5994 | -20.579
12.1536 | -34.5999 | -20.5879
12.1597 | -34.6003 | -20.5968
12.1658 | -34.6008 | -20.6055
12.1719 | -34.6012 | -20.614
12.178 | -34.6017 | -20.6229
12.1841 | -34.6022 | -20.6318
12.1902 | -34.6026 | -20.6406
12.1963 | -34.6031 | -20.6495
12.2024 | -34.6036 | -20.6581
12.2085 | -34.6041 | -20.6669
12.2146 | -34.6045 | -20.6757
12.2207 | -34.605 | -20.6845
12.2268 | -34.6055 | -20.6933
12.2329 | -34.606 | -20.702
12.239 | -34.6065 | -20.7105
12.2451 | -34.6069 | -20.7193
12.2513 | -34.6074 | -20.7282
12.2574 | -34.6078 | -20.737
12.2635 | -34.6083 | -20.7457
12.2696 | -34.6088 | -20.7545
12.2757 | -34.6093 | -20.7629
12.2818 | -34.6098 | -20.7718
12.2879 | -34.6102 | -20.7806
12.294 | -34.6107 | -20.7892
12.3001 | -34.6111 | -20.7979
12.3062 | -34.6116 | -20.8067
12.3123 | -34.6121 | -20.8152
12.3184 | -34.6126 | -20.8238
12.3245 | -34.613 | -20.8324
12.3306 | -34.6135 | -20.8411
12.3367 | -34.614 | -20.8499
12.3428 | -34.6144 | -20.8586
12.3489 | -34.6149 | -20.8671
12.355 | -34.6154 | -20.876
12.3611 | -34.6159 | -20.8848
12.3672 | -34.6163 | -20.8935
12.3733 | -34.6168 | -20.9022
12.3794 | -34.6173 | -20.9107
12.3855 | -34.6177 | -20.9192
12.3916 | -34.6182 | -20.9279
12.3978 | -34.6187 | -20.9367
12.4039 | -34.6192 | -20.9454
12.41 | -34.6197 | -20.9541
12.4161 | -34.6201 | -20.9625
12.4222 | -34.6206 | -20.9712
12.4283 | -34.6211 | -20.9799
12.4344 | -34.6216 | -20.9887
12.4405 | -34.622 | -20.9975
12.4466 | -34.6225 | -21.0061
12.4527 | -34.623 | -21.0146
12.4588 | -34.6234 | -21.0233
12.4649 | -34.6239 | -21.032
12.471 | -34.6244 | -21.0408
12.4771 | -34.6249 | -21.0494
12.4832 | -34.6253 | -21.058
12.4893 | -34.6258 | -21.0664
12.4954 | -34.6263 | -21.075
12.5015 | -34.6267 | -21.0835
12.5076 | -34.6272 | -21.0923
12.5137 | -34.6277 | -21.101
12.5198 | -34.6282 | -21.1095
12.5259 | -34.6286 | -21.1183
12.532 | -34.6291 | -21.127
12.5382 | -34.6296 | -21.1356
12.5443 | -34.6301 | -21.1443
12.5504 | -34.6305 | -21.153
12.5565 | -34.631 | -21.1614
12.5626 | -34.6315 | -21.1702
12.5687 | -34.632 | -21.1789
12.5748 | -34.6324 | -21.1876
12.5809 | -34.6329 | -21.1963
12.587 | -34.6334 | -21.205
12.5931 | -34.6339 | -21.2134
12.5992 | -34.6344 | -21.222
12.6053 | -34.6348 | -21.2306
12.6114 | -34.6353 | -21.2391
12.6175 | -34.6358 | -21.2478
12.6236 | -34.6362 | -21.2564
12.6297 | -34.6367 | -21.265
12.6358 | -34.6372 | -21.2735
12.6419 | -34.6377 | -21.2821
12.648 | -34.6382 | -21.2908
12.6541 | -34.6386 | -21.2994
12.6602 | -34.6391 | -21.3077
12.6663 | -34.6396 | -21.3164
12.6724 | -34.6401 | -21.3248
12.6785 | -34.6405 | -21.3334
12.6847 | -34.641 | -21.3419
12.6908 | -34.6415 | -21.3505
12.6969 | -34.642 | -21.359
12.703 | -34.6425 | -21.3676
12.7091 | -34.6429 | -21.3761
12.7152 | -34.6434 | -21.3845
12.7213 | -34.6439 | -21.3932
12.7274 | -34.6444 | -21.4018
12.7335 | -34.6449 | -21.4101
12.7396 | -34.6453 | -21.4185
12.7457 | -34.6458 | -21.4271
12.7518 | -34.6463 | -21.4356
12.7579 | -34.6468 | -21.4444
12.764 | -34.6472 | -21.4532
12.7701 | -34.6477 | -21.4616
12.7762 | -34.6482 | -21.4701
12.7823 | -34.6487 | -21.4787
12.7884 | -34.6492 | -21.4873
12.7945 | -34.6497 | -21.496
12.8006 | -34.6502 | -21.5044
12.8067 | -34.6507 | -21.513
12.8128 | -34.6511 | -21.5215
12.8189 | -34.6516 | -21.53
12.8251 | -34.6521 | -21.5386
12.8312 | -34.6526 | -21.5474
12.8373 | -34.653 | -21.5558
12.8434 | -34.6535 | -21.5645
12.8495 | -34.654 | -21.573
12.8556 | -34.6545 | -21.5814
12.8617 | -34.6549 | -21.5899
12.8678 | -34.6554 | -21.5986
12.8739 | -34.6559 | -21.607
12.88 | -34.6563 | -21.6155
12.8861 | -34.6568 | -21.624
12.8922 | -34.6573 | -21.6327
12.8983 | -34.6578 | -21.6413
12.9044 | -34.6582 | -21.6497
12.9105 | -34.6587 | -21.6581
12.9166 | -34.6592 | -21.6666
12.9227 | -34.6597 | -21.6751
12.9288 | -34.6602 | -21.6837
12.9349 | -34.6607 | -21.6923
12.941 | -34.6612 | -21.7009
12.9471 | -34.6617 | -21.7095
12.9532 | -34.6622 | -21.718
12.9593 | -34.6627 | -21.7267
12.9654 | -34.6631 | -21.7354
12.9716 | -34.6636 | -21.7441
12.9777 | -34.6641 | -21.7525
12.9838 | -34.6646 | -21.761
12.9899 | -34.665 | -21.7696
12.996 | -34.6655 | -21.7782
13.0021 | -34.666 | -21.7869
13.0082 | -34.6665 | -21.7954
13.0143 | -34.667 | -21.8036
13.0204 | -34.6674 | -21.8121
13.0265 | -34.6679 | -21.8206
13.0326 | -34.6684 | -21.8292
13.0387 | -34.6689 | -21.8379
13.0448 | -34.6693 | -21.8464
13.0509 | -34.6698 | -21.8549
13.057 | -34.6703 | -21.8633
13.0631 | -34.6708 | -21.8718
13.0692 | -34.6713 | -21.8803
13.0753 | -34.6718 | -21.8889
13.0814 | -34.6723 | -21.8975
13.0875 | -34.6728 | -21.9059
13.0936 | -34.6733 | -21.9145
13.0997 | -34.6738 | -21.9228
13.1058 | -34.6742 | -21.9312
13.112 | -34.6747 | -21.9397
13.1181 | -34.6752 | -21.9482
13.1242 | -34.6757 | -21.9567
13.1303 | -34.6762 | -21.9652
13.1364 | -34.6767 | -21.9736
13.1425 | -34.6772 | -21.982
13.1486 | -34.6776 | -21.9906
13.1547 | -34.6781 | -21.999
13.1608 | -34.6786 | -22.0074
13.1669 | -34.6791 | -22.0158
13.173 | -34.6796 | -22.0243
13.1791 | -34.6801 | -22.0328
13.1852 | -34.6805 | -22.0413
13.1913 | -34.681 | -22.0497
13.1974 | -34.6815 | -22.0584
13.2035 | -34.682 | -22.0668
13.2096 | -34.6825 | -22.0752
13.2157 | -34.683 | -22.0834
13.2218 | -34.6835 | -22.092
13.2279 | -34.684 | -22.1005
13.234 | -34.6845 | -22.1087
13.2401 | -34.685 | -22.1171
13.2462 | -34.6855 | -22.1254
13.2524 | -34.6859 | -22.1341
13.2585 | -34.6864 | -22.1425
13.2646 | -34.6869 | -22.1511
13.2707 | -34.6874 | -22.1597
13.2768 | -34.6879 | -22.1682
13.2829 | -34.6884 | -22.1768
13.289 | -34.6888 | -22.1853
13.2951 | -34.6893 | -22.1938
13.3012 | -34.6898 | -22.2024
13.3073 | -34.6903 | -22.2109
13.3134 | -34.6908 | -22.2195
13.3195 | -34.6912 | -22.2279
13.3256 | -34.6917 | -22.2362
13.3317 | -34.6922 | -22.2445
13.3378 | -34.6927 | -22.2532
13.3439 | -34.6932 | -22.2617
13.35 | -34.6937 | -22.2702
13.3561 | -34.6942 | -22.2787
13.3622 | -34.6947 | -22.2869
13.3683 | -34.6952 | -22.2953
13.3744 | -34.6956 | -22.3037
13.3805 | -34.6961 | -22.3123
13.3866 | -34.6966 | -22.3206
13.3927 | -34.6971 | -22.3291
13.3989 | -34.6975 | -22.3374
13.405 | -34.698 | -22.3459
13.4111 | -34.6985 | -22.3543
13.4172 | -34.699 | -22.3627
13.4233 | -34.6995 | -22.3714
13.4294 | -34.6999 | -22.3799
13.4355 | -34.7004 | -22.3882
13.4416 | -34.7009 | -22.3967
13.4477 | -34.7014 | -22.4051
13.4538 | -34.7019 | -22.4136
13.4599 | -34.7024 | -22.4221
13.466 | -34.7028 | -22.4307
13.4721 | -34.7033 | -22.439
13.4782 | -34.7038 | -22.4473
13.4843 | -34.7043 | -22.456
13.4904 | -34.7048 | -22.4645
13.4965 | -34.7053 | -22.4731
13.5026 | -34.7058 | -22.4813
13.5087 | -34.7062 | -22.4894
13.5148 | -34.7067 | -22.498
13.5209 | -34.7071 | -22.5065
13.527 | -34.7076 | -22.5148
13.5331 | -34.7081 | -22.5232
13.5393 | -34.7086 | -22.5316
13.5454 | -34.7091 | -22.54
13.5515 | -34.7096 | -22.5483
13.5576 | -34.7101 | -22.5569
13.5637 | -34.7106 | -22.5656
13.5698 | -34.7111 | -22.574
13.5759 | -34.7115 | -22.5823
13.582 | -34.712 | -22.5906
13.5881 | -34.7125 | -22.5989
13.5942 | -34.713 | -22.6075
13.6003 | -34.7135 | -22.6159
13.6064 | -34.714 | -22.6245
13.6125 | -34.7145 | -22.633
13.6186 | -34.715 | -22.6413
13.6247 | -34.7155 | -22.6499
13.6308 | -34.716 | -22.6584
13.6369 | -34.7165 | -22.6668
13.643 | -34.717 | -22.6754
13.6491 | -34.7175 | -22.6838
13.6552 | -34.718 | -22.6921
13.6613 | -34.7184 | -22.7006
13.6674 | -34.7189 | -22.7091
13.6735 | -34.7194 | -22.7176
13.6796 | -34.7199 | -22.726
13.6858 | -34.7203 | -22.7341
13.6919 | -34.7208 | -22.7425
13.698 | -34.7213 | -22.751
13.7041 | -34.7218 | -22.7597
13.7102 | -34.7223 | -22.7681
13.7163 | -34.7228 | -22.7766
13.7224 | -34.7233 | -22.785
13.7285 | -34.7238 | -22.7934
13.7346 | -34.7242 | -22.8019
13.7407 | -34.7248 | -22.8103
13.7468 | -34.7252 | -22.8187
13.7529 | -34.7257 | -22.8269
13.759 | -34.7262 | -22.8355
13.7651 | -34.7267 | -22.8439
13.7712 | -34.7272 | -22.8524
13.7773 | -34.7276 | -22.8608
13.7834 | -34.7281 | -22.8691
13.7895 | -34.7286 | -22.8776
13.7956 | -34.7291 | -22.886
13.8017 | -34.7295 | -22.8944
13.8078 | -34.73 | -22.9026
13.8139 | -34.7305 | -22.911
13.82 | -34.731 | -22.9193
13.8262 | -34.7314 | -22.9277
13.8323 | -34.7319 | -22.9362
13.8384 | -34.7324 | -22.9446
13.8445 | -34.7329 | -22.953
13.8506 | -34.7334 | -22.9614
13.8567 | -34.7339 | -22.97
13.8628 | -34.7344 | -22.9784
13.8689 | -34.7349 | -22.987
13.875 | -34.7353 | -22.9954
13.8811 | -34.7358 | -23.004
13.8872 | -34.7363 | -23.0125
13.8933 | -34.7368 | -23.0208
13.8994 | -34.7373 | -23.0292
13.9055 | -34.7377 | -23.0377
13.9116 | -34.7382 | -23.0462
13.9177 | -34.7387 | -23.0549
13.9238 | -34.7392 | -23.0632
13.9299 | -34.7397 | -23.0715
13.936 | -34.7401 | -23.08
13.9421 | -34.7406 | -23.0884
13.9482 | -34.7411 | -23.0969
13.9543 | -34.7415 | -23.1056
13.9604 | -34.742 | -23.114
13.9665 | -34.7425 | -23.1225
13.9727 | -34.743 | -23.131
13.9788 | -34.7435 | -23.1393
13.9849 | -34.744 | -23.1477
13.991 | -34.7444 | -23.156
13.9971 | -34.7449 | -23.1643
14.0032 | -34.7454 | -23.1728
14.0093 | -34.7459 | -23.1812
14.0154 | -34.7463 | -23.1896
14.0215 | -34.7468 | -23.198
14.0276 | -34.7473 | -23.2064
14.0337 | -34.7477 | -23.2147
14.0398 | -34.7482 | -23.223
14.0459 | -34.7487 | -23.2314
14.052 | -34.7492 | -23.2398
14.0581 | -34.7496 | -23.2481
14.0642 | -34.7501 | -23.2564
14.0703 | -34.7506 | -23.2647
14.0764 | -34.7511 | -23.2729
14.0825 | -34.7516 | -23.2812
14.0886 | -34.7521 | -23.2895
14.0947 | -34.7526 | -23.2979
14.1008 | -34.753 | -23.3063
14.1069 | -34.7535 | -23.3145
14.1131 | -34.7539 | -23.323
14.1192 | -34.7544 | -23.3314
14.1253 | -34.7549 | -23.3397
14.1314 | -34.7554 | -23.348
14.1375 | -34.7559 | -23.3563
14.1436 | -34.7564 | -23.3646
14.1497 | -34.7568 | -23.3729
14.1558 | -34.7573 | -23.3813
14.1619 | -34.7578 | -23.3895
14.168 | -34.7583 | -23.3978
14.1741 | -34.7588 | -23.4061
14.1802 | -34.7593 | -23.4142
14.1863 | -34.7597 | -23.4226
14.1924 | -34.7603 | -23.431
14.1985 | -34.7608 | -23.4392
14.2046 | -34.7612 | -23.4478
14.2107 | -34.7618 | -23.4559
14.2168 | -34.7623 | -23.4644
14.2229 | -34.7627 | -23.4729
14.229 | -34.7632 | -23.4813
14.2351 | -34.7637 | -23.4896
14.2412 | -34.7642 | -23.498
14.2473 | -34.7647 | -23.5062
14.2534 | -34.7652 | -23.5145
14.2596 | -34.7656 | -23.5229
14.2657 | -34.7661 | -23.5311
14.2718 | -34.7666 | -23.5395
14.2779 | -34.7671 | -23.5479
14.284 | -34.7676 | -23.5565
14.2901 | -34.7681 | -23.5649
14.2962 | -34.7685 | -23.5731
14.3023 | -34.769 | -23.5814
14.3084 | -34.7695 | -23.5895
14.3145 | -34.77 | -23.5979
14.3206 | -34.7704 | -23.6063
14.3267 | -34.7709 | -23.6147
14.3328 | -34.7714 | -23.6229
14.3389 | -34.7719 | -23.6312
14.345 | -34.7724 | -23.6396
14.3511 | -34.7729 | -23.6478
14.3572 | -34.7734 | -23.6561
14.3633 | -34.7738 | -23.6647
14.3694 | -34.7743 | -23.6728
14.3755 | -34.7749 | -23.6812
14.3816 | -34.7754 | -23.6896
14.3877 | -34.7758 | -23.6978
14.3938 | -34.7763 | -23.7062
14.4 | -34.7768 | -23.7146
14.4061 | -34.7773 | -23.7231
14.4122 | -34.7778 | -23.7314
14.4183 | -34.7782 | -23.7398
14.4244 | -34.7787 | -23.7483
14.4305 | -34.7792 | -23.7563
14.4366 | -34.7797 | -23.7647
14.4427 | -34.7802 | -23.7733
14.4488 | -34.7807 | -23.7817
14.4549 | -34.7812 | -23.7899
14.461 | -34.7816 | -23.798
14.4671 | -34.7821 | -23.8066
14.4732 | -34.7826 | -23.8148
14.4793 | -34.7831 | -23.8231
14.4854 | -34.7836 | -23.8313
14.4915 | -34.7841 | -23.8395
14.4976 | -34.7845 | -23.8478
14.5037 | -34.7851 | -23.8562
14.5098 | -34.7855 | -23.8645
14.5159 | -34.786 | -23.8728
14.522 | -34.7865 | -23.8812
14.5281 | -34.787 | -23.8894
14.5342 | -34.7875 | -23.8978
14.5403 | -34.788 | -23.9061
14.5465 | -34.7884 | -23.9145
14.5526 | -34.7889 | -23.9231
14.5587 | -34.7894 | -23.9316
14.5648 | -34.7899 | -23.9399
14.5709 | -34.7904 | -23.9481
14.577 | -34.7909 | -23.9563
14.5831 | -34.7913 | -23.9646
14.5892 | -34.7918 | -23.9726
14.5953 | -34.7923 | -23.9808
14.6014 | -34.7928 | -23.9892
14.6075 | -34.7933 | -23.9974
14.6136 | -34.7937 | -24.0056
14.6197 | -34.7942 | -24.0138
14.6258 | -34.7947 | -24.0221
14.6319 | -34.7952 | -24.0303
14.638 | -34.7957 | -24.0386
14.6441 | -34.7962 | -24.0472
14.6502 | -34.7967 | -24.0555
14.6563 | -34.7972 | -24.0636
14.6624 | -34.7976 | -24.072
14.6685 | -34.7982 | -24.0799
14.6746 | -34.7986 | -24.0882
14.6807 | -34.7991 | -24.0961
14.6869 | -34.7996 | -24.1041
14.693 | -34.8001 | -24.1121
14.6991 | -34.8006 | -24.1202
14.7052 | -34.8011 | -24.1283
14.7113 | -34.8015 | -24.1368
14.7174 | -34.802 | -24.145
14.7235 | -34.8025 | -24.1532
14.7296 | -34.803 | -24.1615
14.7357 | -34.8035 | -24.1696
14.7418 | -34.804 | -24.1778
14.7479 | -34.8044 | -24.1861
14.754 | -34.8049 | -24.1945
14.7601 | -34.8054 | -24.2025
14.7662 | -34.8059 | -24.2106
14.7723 | -34.8063 | -24.2184
14.7784 | -34.8068 | -24.2268
14.7845 | -34.8073 | -24.2349
14.7906 | -34.8078 | -24.2433
14.7967 | -34.8083 | -24.2515
14.8028 | -34.8088 | -24.2597
14.8089 | -34.8093 | -24.2676
14.815 | -34.8098 | -24.2758
14.8211 | -34.8102 | -24.2839
14.8272 | -34.8107 | -24.2922
14.8334 | -34.8112 | -24.3003
14.8395 | -34.8116 | -24.3083
14.8456 | -34.8121 | -24.3164
14.8517 | -34.8126 | -24.3246
14.8578 | -34.8131 | -24.3329
14.8639 | -34.8136 | -24.3412
14.87 | -34.8141 | -24.3496
14.8761 | -34.8145 | -24.3579
14.8822 | -34.815 | -24.3662
14.8883 | -34.8154 | -24.3744
14.8944 | -34.8159 | -24.3825
14.9005 | -34.8164 | -24.3909
14.9066 | -34.8169 | -24.3991
14.9127 | -34.8174 | -24.407
14.9188 | -34.8179 | -24.4153
14.9249 | -34.8183 | -24.4234
14.931 | -34.8188 | -24.4317
14.9371 | -34.8193 | -24.4397
14.9432 | -34.8199 | -24.4479
14.9493 | -34.8204 | -24.4561
14.9554 | -34.8209 | -24.4643
14.9615 | -34.8213 | -24.4723
14.9676 | -34.8218 | -24.4803
14.9738 | -34.8223 | -24.4886
14.9799 | -34.8227 | -24.4968
14.986 | -34.8232 | -24.505
14.9921 | -34.8238 | -24.5132
14.9982 | -34.8242 | -24.5214
15.0043 | -34.8246 | -24.5293
15.0104 | -34.8251 | -24.5374
15.0165 | -34.8256 | -24.5457
15.0226 | -34.8262 | -24.5535
15.0287 | -34.8267 | -24.5617
15.0348 | -34.8271 | -24.5701
15.0409 | -34.8276 | -24.5784
15.047 | -34.8281 | -24.5867
15.0531 | -34.8286 | -24.5946
15.0592 | -34.8291 | -24.6029
15.0653 | -34.8296 | -24.6111
15.0714 | -34.8301 | -24.6191
15.0775 | -34.8306 | -24.6273
15.0836 | -34.8311 | -24.6355
15.0897 | -34.8317 | -24.6434
15.0958 | -34.8322 | -24.6515
15.1019 | -34.8326 | -24.6595
15.108 | -34.8331 | -24.6676
15.1141 | -34.8336 | -24.6756
15.1203 | -34.8341 | -24.6839
15.1264 | -34.8346 | -24.692
15.1325 | -34.8351 | -24.7002
15.1386 | -34.8356 | -24.7083
15.1447 | -34.8361 | -24.7164
15.1508 | -34.8366 | -24.7245
15.1569 | -34.8371 | -24.7325
15.163 | -34.8376 | -24.7402
15.1691 | -34.8381 | -24.7481
15.1752 | -34.8385 | -24.756
15.1813 | -34.839 | -24.7642
15.1874 | -34.8395 | -24.7722
15.1935 | -34.84 | -24.7801
15.1996 | -34.8405 | -24.7881
15.2057 | -34.841 | -24.7961
15.2118 | -34.8414 | -24.8043
15.2179 | -34.8419 | -24.8125
15.224 | -34.8424 | -24.8206
15.2301 | -34.8429 | -24.8285
15.2362 | -34.8434 | -24.8365
15.2423 | -34.8438 | -24.8447
15.2484 | -34.8443 | -24.8529
15.2545 | -34.8448 | -24.8611
15.2607 | -34.8453 | -24.8691
15.2668 | -34.8458 | -24.8767
15.2729 | -34.8463 | -24.8846
15.279 | -34.8468 | -24.8928
15.2851 | -34.8473 | -24.9007
15.2912 | -34.8478 | -24.909
15.2973 | -34.8483 | -24.9171
15.3034 | -34.8488 | -24.9249
15.3095 | -34.8494 | -24.9328
15.3156 | -34.8498 | -24.9409
15.3217 | -34.8503 | -24.9488
15.3278 | -34.8508 | -24.957
15.3339 | -34.8513 | -24.9647
15.34 | -34.8518 | -24.9727
15.3461 | -34.8523 | -24.9806
15.3522 | -34.8528 | -24.9886
15.3583 | -34.8533 | -24.9967
15.3644 | -34.8538 | -25.0047
15.3705 | -34.8543 | -25.0124
15.3766 | -34.8548 | -25.0206
15.3827 | -34.8553 | -25.0287
15.3888 | -34.8558 | -25.0369
15.3949 | -34.8563 | -25.0449
15.401 | -34.8567 | -25.0531
15.4072 | -34.8572 | -25.0607
15.4133 | -34.8577 | -25.0686
15.4194 | -34.8582 | -25.0768
15.4255 | -34.8587 | -25.0849
15.4316 | -34.8592 | -25.093
15.4377 | -34.8597 | -25.101
15.4438 | -34.8602 | -25.109
15.4499 | -34.8607 | -25.117
15.456 | -34.8612 | -25.1249
15.4621 | -34.8617 | -25.1329
15.4682 | -34.8622 | -25.141
15.4743 | -34.8627 | -25.1488
15.4804 | -34.8631 | -25.1568
15.4865 | -34.8636 | -25.1648
15.4926 | -34.8641 | -25.1729
15.4987 | -34.8646 | -25.1809
15.5048 | -34.8651 | -25.189
15.5109 | -34.8656 | -25.1968
15.517 | -34.8661 | -25.2049
15.5231 | -34.8666 | -25.2129
15.5292 | -34.8671 | -25.2208
15.5353 | -34.8676 | -25.2289
15.5414 | -34.8681 | -25.2368
15.5476 | -34.8686 | -25.2444
15.5537 | -34.8691 | -25.2523
15.5598 | -34.8696 | -25.2602
15.5659 | -34.8701 | -25.2681
15.572 | -34.8706 | -25.2761
15.5781 | -34.871 | -25.2838
15.5842 | -34.8716 | -25.2916
15.5903 | -34.8721 | -25.2994
15.5964 | -34.8726 | -25.3072
15.6025 | -34.873 | -25.3151
15.6086 | -34.8736 | -25.3231
15.6147 | -34.874 | -25.3311
15.6208 | -34.8746 | -25.3388
15.6269 | -34.8751 | -25.3467
15.633 | -34.8756 | -25.3547
15.6391 | -34.8761 | -25.3626
15.6452 | -34.8766 | -25.3706
15.6513 | -34.8771 | -25.3782
15.6574 | -34.8776 | -25.3862
15.6635 | -34.8781 | -25.3941
15.6696 | -34.8786 | -25.402
15.6757 | -34.8791 | -25.4101
15.6818 | -34.8796 | -25.418
15.688 | -34.8801 | -25.4258
15.6941 | -34.8806 | -25.4337
15.7002 | -34.8811 | -25.4416
15.7063 | -34.8816 | -25.4495
15.7124 | -34.882 | -25.4575
15.7185 | -34.8825 | -25.4654
15.7246 | -34.8831 | -25.4733
15.7307 | -34.8836 | -25.4811
15.7368 | -34.8841 | -25.4889
15.7429 | -34.8846 | -25.4968
15.749 | -34.8851 | -25.5049
15.7551 | -34.8856 | -25.5129
15.7612 | -34.8861 | -25.5207
15.7673 | -34.8866 | -25.5286
15.7734 | -34.8871 | -25.5363
15.7795 | -34.8876 | -25.5443
15.7856 | -34.8881 | -25.552
15.7917 | -34.8886 | -25.5596
15.7978 | -34.8891 | -25.5676
15.8039 | -34.8896 | -25.5755
15.81 | -34.8901 | -25.5835
15.8161 | -34.8906 | -25.5912
15.8222 | -34.8911 | -25.599
15.8283 | -34.8916 | -25.6066
15.8345 | -34.8921 | -25.6145
15.8406 | -34.8926 | -25.6223
15.8467 | -34.8931 | -25.6303
15.8528 | -34.8937 | -25.6381
15.8589 | -34.8942 | -25.646
15.865 | -34.8947 | -25.6535
15.8711 | -34.8952 | -25.6614
15.8772 | -34.8957 | -25.6691
15.8833 | -34.8962 | -25.6771
15.8894 | -34.8968 | -25.685
15.8955 | -34.8972 | -25.6927
15.9016 | -34.8977 | -25.7003
15.9077 | -34.8982 | -25.7081
15.9138 | -34.8987 | -25.7159
15.9199 | -34.8992 | -25.7236
15.926 | -34.8997 | -25.7314
15.9321 | -34.9002 | -25.7391
15.9382 | -34.9007 | -25.7467
15.9443 | -34.9012 | -25.7547
15.9504 | -34.9017 | -25.7626
15.9565 | -34.9022 | -25.7704
15.9626 | -34.9027 | -25.7782
15.9687 | -34.9032 | -25.7858
15.9749 | -34.9037 | -25.7936
15.981 | -34.9042 | -25.8014
15.9871 | -34.9047 | -25.8094
15.9932 | -34.9052 | -25.8169
15.9993 | -34.9057 | -25.8247
16.0054 | -34.9063 | -25.8321
16.0115 | -34.9068 | -25.8401
16.0176 | -34.9073 | -25.8479
16.0237 | -34.9078 | -25.8556
16.0298 | -34.9083 | -25.8633
16.0359 | -34.9088 | -25.8709
16.042 | -34.9093 | -25.8787
16.0481 | -34.9098 | -25.8866
16.0542 | -34.9103 | -25.8945
16.0603 | -34.9108 | -25.902
16.0664 | -34.9113 | -25.9098
16.0725 | -34.9118 | -25.9174
16.0786 | -34.9123 | -25.925
16.0847 | -34.9128 | -25.9328
16.0908 | -34.9133 | -25.9407
16.0969 | -34.9138 | -25.9485
16.103 | -34.9143 | -25.9563
16.1091 | -34.9148 | -25.964
16.1152 | -34.9153 | -25.9717
16.1214 | -34.9158 | -25.9795
16.1275 | -34.9163 | -25.9873
16.1336 | -34.9168 | -25.9949
16.1397 | -34.9173 | -26.0025
16.1458 | -34.9179 | -26.0101
16.1519 | -34.9184 | -26.018
16.158 | -34.9189 | -26.0258
16.1641 | -34.9194 | -26.0337
16.1702 | -34.9199 | -26.0415
16.1763 | -34.9204 | -26.0493
16.1824 | -34.9209 | -26.0569
16.1885 | -34.9214 | -26.0645
16.1946 | -34.9219 | -26.0722
16.2007 | -34.9224 | -26.08
16.2068 | -34.9229 | -26.0877
16.2129 | -34.9234 | -26.0955
16.219 | -34.9239 | -26.1029
16.2251 | -34.9244 | -26.1105
16.2312 | -34.925 | -26.1183
16.2373 | -34.9255 | -26.1259
16.2434 | -34.926 | -26.1336
16.2495 | -34.9265 | -26.1411
16.2556 | -34.927 | -26.1489
16.2618 | -34.9275 | -26.1566
16.2679 | -34.928 | -26.1643
16.274 | -34.9285 | -26.172
16.2801 | -34.929 | -26.1797
16.2862 | -34.9296 | -26.1873
16.2923 | -34.9301 | -26.1951
16.2984 | -34.9306 | -26.2029
16.3045 | -34.9311 | -26.2107
16.3106 | -34.9316 | -26.2185
16.3167 | -34.9321 | -26.2262
16.3228 | -34.9326 | -26.2337
16.3289 | -34.9331 | -26.2414
16.335 | -34.9336 | -26.249
16.3411 | -34.9341 | -26.2568
16.3472 | -34.9346 | -26.2646
16.3533 | -34.9352 | -26.2723
16.3594 | -34.9357 | -26.2797
16.3655 | -34.9362 | -26.2874
16.3716 | -34.9367 | -26.2951
16.3777 | -34.9372 | -26.3028
16.3838 | -34.9377 | -26.3105
16.3899 | -34.9382 | -26.3181
16.396 | -34.9387 | -26.3258
16.4021 | -34.9392 | -26.3333
16.4083 | -34.9397 | -26.341
16.4144 | -34.9402 | -26.3487
16.4205 | -34.9407 | -26.3563
16.4266 | -34.9412 | -26.3638
16.4327 | -34.9418 | -26.3715
16.4388 | -34.9423 | -26.3793
16.4449 | -34.9428 | -26.3869
16.451 | -34.9433 | -26.3947
16.4571 | -34.9438 | -26.4026
16.4632 | -34.9443 | -26.41
16.4693 | -34.9448 | -26.4178
16.4754 | -34.9453 | -26.4257
16.4815 | -34.9458 | -26.4334
16.4876 | -34.9463 | -26.441
16.4937 | -34.9468 | -26.4486
16.4998 | -34.9473 | -26.4562
16.5059 | -34.9478 | -26.464
16.512 | -34.9483 | -26.4717
16.5181 | -34.9488 | -26.4795
16.5242 | -34.9493 | -26.4872
16.5303 | -34.9498 | -26.4948
16.5364 | -34.9503 | -26.5024
16.5425 | -34.9508 | -26.5101
16.5487 | -34.9512 | -26.5177
16.5548 | -34.9517 | -26.5254
16.5609 | -34.9523 | -26.533
16.567 | -34.9528 | -26.5406
16.5731 | -34.9533 | -26.5483
16.5792 | -34.9538 | -26.556
16.5853 | -34.9543 | -26.5637
16.5914 | -34.9548 | -26.5714
16.5975 | -34.9553 | -26.579
16.6036 | -34.9558 | -26.5865
16.6097 | -34.9563 | -26.5943
16.6158 | -34.9568 | -26.6018
16.6219 | -34.9573 | -26.6095
16.628 | -34.9578 | -26.6171
16.6341 | -34.9584 | -26.6248
16.6402 | -34.9589 | -26.6323
16.6463 | -34.9593 | -26.6399
16.6524 | -34.9598 | -26.6474
16.6585 | -34.9603 | -26.6552
16.6646 | -34.9609 | -26.6628
16.6707 | -34.9614 | -26.6703
16.6768 | -34.9619 | -26.678
16.6829 | -34.9624 | -26.6855
16.689 | -34.9629 | -26.6931
16.6952 | -34.9634 | -26.7008
16.7013 | -34.9639 | -26.7084
16.7074 | -34.9644 | -26.7159
16.7135 | -34.9649 | -26.7235
16.7196 | -34.9654 | -26.7311
16.7257 | -34.9659 | -26.7388
16.7318 | -34.9664 | -26.7463
16.7379 | -34.9669 | -26.754
16.744 | -34.9674 | -26.7614
16.7501 | -34.9679 | -26.7691
16.7562 | -34.9684 | -26.7767
16.7623 | -34.9689 | -26.7841
16.7684 | -34.9694 | -26.7919
16.7745 | -34.9699 | -26.7997
16.7806 | -34.9704 | -26.807
16.7867 | -34.971 | -26.8147
16.7928 | -34.9715 | -26.8223
16.7989 | -34.972 | -26.8299
16.805 | -34.9725 | -26.8376
16.8111 | -34.973 | -26.845
16.8172 | -34.9735 | -26.8527
16.8233 | -34.974 | -26.8604
16.8294 | -34.9745 | -26.868
16.8356 | -34.975 | -26.8758
16.8417 | -34.9755 | -26.8833
16.8478 | -34.976 | -26.8909
16.8539 | -34.9765 | -26.8985
16.86 | -34.977 | -26.9063
16.8661 | -34.9775 | -26.9138
16.8722 | -34.978 | -26.9215
16.8783 | -34.9785 | -26.9292
16.8844 | -34.9791 | -26.9366
16.8905 | -34.9796 | -26.9442
16.8966 | -34.9801 | -26.9518
16.9027 | -34.9806 | -26.9595
16.9088 | -34.981 | -26.9672
16.9149 | -34.9815 | -26.9749
16.921 | -34.9821 | -26.9823
16.9271 | -34.9826 | -26.99
16.9332 | -34.9831 | -26.9976
16.9393 | -34.9836 | -27.0052
16.9454 | -34.9841 | -27.0127
16.9515 | -34.9846 | -27.0204
16.9576 | -34.9851 | -27.0279
16.9637 | -34.9856 | -27.0355
16.9698 | -34.9861 | -27.043
16.9759 | -34.9866 | -27.0505
16.9821 | -34.9871 | -27.0581
16.9882 | -34.9877 | -27.0655
16.9943 | -34.9882 | -27.073
17.0004 | -34.9887 | -27.0805
17.0065 | -34.9892 | -27.0881
17.0126 | -34.9897 | -27.0958
17.0187 | -34.9902 | -27.1035
17.0248 | -34.9907 | -27.1109
17.0309 | -34.9912 | -27.1184
17.037 | -34.9917 | -27.1259
17.0431 | -34.9922 | -27.1334
17.0492 | -34.9927 | -27.1411
17.0553 | -34.9932 | -27.1487
17.0614 | -34.9937 | -27.1563
17.0675 | -34.9942 | -27.1638
17.0736 | -34.9947 | -27.1714
17.0797 | -34.9952 | -27.1789
17.0858 | -34.9957 | -27.1865
17.0919 | -34.9962 | -27.1941
17.098 | -34.9967 | -27.2016
17.1041 | -34.9972 | -27.2091
17.1102 | -34.9977 | -27.2166
17.1163 | -34.9982 | -27.2241
17.1225 | -34.9987 | -27.2316
17.1286 | -34.9992 | -27.2391
17.1347 | -34.9998 | -27.2466
17.1408 | -35.0003 | -27.2544
17.1469 | -35.0008 | -27.2619
17.153 | -35.0013 | -27.2694
17.1591 | -35.0018 | -27.277
17.1652 | -35.0023 | -27.2843
17.1713 | -35.0028 | -27.2918
17.1774 | -35.0033 | -27.2994
17.1835 | -35.0038 | -27.3069
17.1896 | -35.0043 | -27.3145
17.1957 | -35.0048 | -27.322
17.2018 | -35.0053 | -27.3294
17.2079 | -35.0059 | -27.337
17.214 | -35.0064 | -27.3447
17.2201 | -35.0069 | -27.3522
17.2262 | -35.0074 | -27.3599
17.2323 | -35.0079 | -27.3675
17.2384 | -35.0084 | -27.375
17.2445 | -35.0089 | -27.3824
17.2506 | -35.0094 | -27.39
17.2567 | -35.0099 | -27.3976
17.2628 | -35.0104 | -27.4051
17.269 | -35.0109 | -27.4126
17.2751 | -35.0114 | -27.4201
17.2812 | -35.0119 | -27.4278
17.2873 | -35.0124 | -27.4354
17.2934 | -35.0129 | -27.4431
17.2995 | -35.0134 | -27.4508
17.3056 | -35.0139 | -27.4582
17.3117 | -35.0144 | -27.4659
17.3178 | -35.0149 | -27.4733
17.3239 | -35.0154 | -27.4808
17.33 | -35.0159 | -27.4884
17.3361 | -35.0164 | -27.496
17.3422 | -35.0169 | -27.5033
17.3483 | -35.0174 | -27.5109
17.3544 | -35.0179 | -27.5185
17.3605 | -35.0184 | -27.526
17.3666 | -35.0189 | -27.5335
17.3727 | -35.0194 | -27.541
17.3788 | -35.0199 | -27.5483
17.3849 | -35.0204 | -27.5558
17.391 | -35.0209 | -27.5634
17.3971 | -35.0214 | -27.5708
17.4032 | -35.0219 | -27.5783
17.4094 | -35.0224 | -27.5857
17.4155 | -35.0229 | -27.5932
17.4216 | -35.0234 | -27.6007
17.4277 | -35.0239 | -27.6083
17.4338 | -35.0244 | -27.6158
17.4399 | -35.0249 | -27.6234
17.446 | -35.0254 | -27.6308
17.4521 | -35.0259 | -27.6383
17.4582 | -35.0264 | -27.6457
17.4643 | -35.0269 | -27.6531
17.4704 | -35.0274 | -27.6607
17.4765 | -35.0279 | -27.6683
17.4826 | -35.0284 | -27.6758
17.4887 | -35.0289 | -27.6833
17.4948 | -35.0294 | -27.6908
17.5009 | -35.0299 | -27.6983
17.507 | -35.0304 | -27.7059
17.5131 | -35.0309 | -27.7133
17.5192 | -35.0314 | -27.7208
17.5253 | -35.0319 | -27.7284
17.5314 | -35.0323 | -27.736
17.5375 | -35.0328 | -27.7436
17.5436 | -35.0333 | -27.7512
17.5497 | -35.0338 | -27.7587
17.5559 | -35.0343 | -27.766
17.562 | -35.0349 | -27.7735
17.5681 | -35.0354 | -27.781
17.5742 | -35.0359 | -27.7887
17.5803 | -35.0364 | -27.7961
17.5864 | -35.0369 | -27.8035
17.5925 | -35.0374 | -27.8111
17.5986 | -35.0379 | -27.8186
17.6047 | -35.0384 | -27.8262
17.6108 | -35.0389 | -27.8336
17.6169 | -35.0394 | -27.8412
17.623 | -35.0399 | -27.8485
17.6291 | -35.0404 | -27.856
17.6352 | -35.0409 | -27.8635
17.6413 | -35.0414 | -27.8711
17.6474 | -35.0419 | -27.8786
17.6535 | -35.0424 | -27.8861
17.6596 | -35.0429 | -27.8935
17.6657 | -35.0434 | -27.901
17.6718 | -35.0439 | -27.9085
17.6779 | -35.0444 | -27.9159
17.684 | -35.0449 | -27.9235
17.6901 | -35.0454 | -27.931
17.6963 | -35.0459 | -27.9383
17.7024 | -35.0464 | -27.9458
17.7085 | -35.0469 | -27.9532
17.7146 | -35.0474 | -27.9606
17.7207 | -35.0479 | -27.968
17.7268 | -35.0485 | -27.9754
17.7329 | -35.049 | -27.9829
17.739 | -35.0494 | -27.9905
17.7451 | -35.05 | -27.9979
17.7512 | -35.0505 | -28.0055
17.7573 | -35.051 | -28.0132
17.7634 | -35.0515 | -28.0205
17.7695 | -35.052 | -28.0281
17.7756 | -35.0525 | -28.0357
17.7817 | -35.053 | -28.0433
17.7878 | -35.0535 | -28.0508
17.7939 | -35.054 | -28.0583
17.8 | -35.0545 | -28.0656
17.8061 | -35.055 | -28.0731
17.8122 | -35.0555 | -28.0806
17.8183 | -35.056 | -28.0881
17.8244 | -35.0565 | -28.0956
17.8305 | -35.057 | -28.103
17.8366 | -35.0576 | -28.1103
17.8428 | -35.058 | -28.1178
17.8489 | -35.0585 | -28.1254
17.855 | -35.059 | -28.1327
17.8611 | -35.0595 | -28.1401
17.8672 | -35.0601 | -28.1474
17.8733 | -35.0606 | -28.155
17.8794 | -35.0611 | -28.1624
17.8855 | -35.0616 | -28.17
17.8916 | -35.0621 | -28.1775
17.8977 | -35.0625 | -28.1851
17.9038 | -35.0631 | -28.1925
17.9099 | -35.0636 | -28.1999
17.916 | -35.0641 | -28.2074
17.9221 | -35.0646 | -28.215
17.9282 | -35.0651 | -28.2224
17.9343 | -35.0656 | -28.23
17.9404 | -35.0661 | -28.2373
17.9465 | -35.0666 | -28.2448
17.9526 | -35.0671 | -28.2523
17.9587 | -35.0676 | -28.2597
17.9648 | -35.0681 | -28.2669
17.9709 | -35.0685 | -28.2744
17.977 | -35.069 | -28.2818
17.9832 | -35.0695 | -28.2893
17.9893 | -35.07 | -28.2968
17.9954 | -35.0706 | -28.3043
18.0015 | -35.0711 | -28.3119
18.0076 | -35.0716 | -28.3191
18.0137 | -35.072 | -28.3266
18.0198 | -35.0725 | -28.334
18.0259 | -35.073 | -28.3416
18.032 | -35.0735 | -28.349
18.0381 | -35.074 | -28.3565
18.0442 | -35.0745 | -28.3639
18.0503 | -35.075 | -28.3713
18.0564 | -35.0755 | -28.3788
18.0625 | -35.0759 | -28.3864
18.0686 | -35.0764 | -28.3938
18.0747 | -35.0769 | -28.4013
18.0808 | -35.0775 | -28.4085
18.0869 | -35.078 | -28.416
18.093 | -35.0785 | -28.4234
18.0991 | -35.079 | -28.431
18.1052 | -35.0795 | -28.4385
18.1113 | -35.08 | -28.4459
18.1174 | -35.0805 | -28.4532
18.1236 | -35.081 | -28.4606
18.1297 | -35.0815 | -28.468
18.1358 | -35.082 | -28.4754
18.1419 | -35.0825 | -28.4828
18.148 | -35.083 | -28.4902
18.1541 | -35.0835 | -28.4974
18.1602 | -35.084 | -28.5049
18.1663 | -35.0845 | -28.5123
18.1724 | -35.085 | -28.5198
18.1785 | -35.0854 | -28.5272
18.1846 | -35.0859 | -28.5345
18.1907 | -35.0864 | -28.542
18.1968 | -35.0869 | -28.5494
18.2029 | -35.0874 | -28.5567
18.209 | -35.0879 | -28.5642
18.2151 | -35.0884 | -28.5718
18.2212 | -35.0889 | -28.5791
18.2273 | -35.0894 | -28.5866
18.2334 | -35.0899 | -28.594
18.2395 | -35.0904 | -28.6013
18.2456 | -35.0909 | -28.6088
18.2517 | -35.0914 | -28.6162
18.2578 | -35.0919 | -28.6236
18.2639 | -35.0924 | -28.6311
18.2701 | -35.0929 | -28.6384
18.2762 | -35.0934 | -28.6458
18.2823 | -35.0939 | -28.6532
18.2884 | -35.0945 | -28.6605
18.2945 | -35.095 | -28.6676
18.3006 | -35.0955 | -28.675
18.3067 | -35.0959 | -28.6824
18.3128 | -35.0964 | -28.6899
18.3189 | -35.0969 | -28.6973
18.325 | -35.0974 | -28.7044
18.3311 | -35.0979 | -28.7118
18.3372 | -35.0984 | -28.7192
18.3433 | -35.0989 | -28.7267
18.3494 | -35.0994 | -28.7342
18.3555 | -35.0999 | -28.7416
18.3616 | -35.1004 | -28.7489
18.3677 | -35.1009 | -28.7562
18.3738 | -35.1014 | -28.7635
18.3799 | -35.1019 | -28.7708
18.386 | -35.1024 | -28.7781
18.3921 | -35.1029 | -28.7854
18.3982 | -35.1034 | -28.7926
18.4043 | -35.1038 | -28.7998
18.4105 | -35.1043 | -28.8072
18.4166 | -35.1048 | -28.8145
18.4227 | -35.1053 | -28.8218
18.4288 | -35.1058 | -28.8292
18.4349 | -35.1063 | -28.8363
18.441 | -35.1068 | -28.8436
18.4471 | -35.1073 | -28.851
18.4532 | -35.1078 | -28.8583
18.4593 | -35.1083 | -28.8656
18.4654 | -35.1088 | -28.8728
18.4715 | -35.1093 | -28.8803
18.4776 | -35.1098 | -28.8877
18.4837 | -35.1103 | -28.8951
18.4898 | -35.1108 | -28.9025
18.4959 | -35.1113 | -28.9099
18.502 | -35.1117 | -28.9172
18.5081 | -35.1122 | -28.9245
18.5142 | -35.1127 | -28.9319
18.5203 | -35.1132 | -28.9392
18.5264 | -35.1137 | -28.9465
18.5325 | -35.1142 | -28.9539
18.5386 | -35.1147 | -28.9611
18.5447 | -35.1152 | -28.9683
18.5508 | -35.1157 | -28.9756
18.557 | -35.1162 | -28.983
18.5631 | -35.1167 | -28.9903
18.5692 | -35.1172 | -28.9976
18.5753 | -35.1177 | -29.0048
18.5814 | -35.1182 | -29.0121
18.5875 | -35.1187 | -29.0194
18.5936 | -35.1191 | -29.0265
18.5997 | -35.1196 | -29.034
18.6058 | -35.1201 | -29.0412
18.6119 | -35.1206 | -29.0486
18.618 | -35.1211 | -29.0559
18.6241 | -35.1216 | -29.0631
18.6302 | -35.1221 | -29.0706
18.6363 | -35.1226 | -29.078
18.6424 | -35.123 | -29.0852
18.6485 | -35.1235 | -29.0923
18.6546 | -35.124 | -29.0995
18.6607 | -35.1245 | -29.1069
18.6668 | -35.125 | -29.1142
18.6729 | -35.1255 | -29.1216
18.679 | -35.126 | -29.1288
18.6851 | -35.1265 | -29.1361
18.6912 | -35.127 | -29.1433
18.6974 | -35.1274 | -29.1506
18.7035 | -35.1279 | -29.158
18.7096 | -35.1284 | -29.1654
18.7157 | -35.1289 | -29.1725
18.7218 | -35.1294 | -29.1798
18.7279 | -35.1299 | -29.1871
18.734 | -35.1304 | -29.1945
18.7401 | -35.1309 | -29.2017
18.7462 | -35.1314 | -29.2088
18.7523 | -35.1319 | -29.2161
18.7584 | -35.1324 | -29.2235
18.7645 | -35.1329 | -29.2307
18.7706 | -35.1334 | -29.238
18.7767 | -35.1339 | -29.2454
18.7828 | -35.1344 | -29.2524
18.7889 | -35.1349 | -29.2596
18.795 | -35.1354 | -29.267
18.8011 | -35.1359 | -29.2743
18.8072 | -35.1363 | -29.2818
18.8133 | -35.1368 | -29.2892
18.8194 | -35.1373 | -29.2961
18.8255 | -35.1378 | -29.3034
18.8316 | -35.1383 | -29.3107
18.8377 | -35.1388 | -29.3178
18.8439 | -35.1393 | -29.325
18.85 | -35.1397 | -29.3321
18.8561 | -35.1402 | -29.3392
18.8622 | -35.1407 | -29.3464
18.8683 | -35.1412 | -29.3538
18.8744 | -35.1417 | -29.361
18.8805 | -35.1422 | -29.3683
18.8866 | -35.1427 | -29.3757
18.8927 | -35.1432 | -29.3828
18.8988 | -35.1437 | -29.39
18.9049 | -35.1442 | -29.3971
18.911 | -35.1447 | -29.4043
18.9171 | -35.1452 | -29.4115
18.9232 | -35.1457 | -29.4186
18.9293 | -35.1462 | -29.4259
18.9354 | -35.1467 | -29.4332
18.9415 | -35.1472 | -29.4405
18.9476 | -35.1476 | -29.4476
18.9537 | -35.1481 | -29.4548
18.9598 | -35.1486 | -29.4618
18.9659 | -35.1491 | -29.469
18.972 | -35.1496 | -29.4763
18.9781 | -35.1501 | -29.4835
18.9843 | -35.1505 | -29.4909
18.9904 | -35.151 | -29.4981
18.9965 | -35.1515 | -29.5051
19.0026 | -35.152 | -29.5125
19.0087 | -35.1525 | -29.5197
19.0148 | -35.153 | -29.5267
19.0209 | -35.1535 | -29.5341
19.027 | -35.154 | -29.5413
19.0331 | -35.1545 | -29.5485
19.0392 | -35.155 | -29.5558
19.0453 | -35.1555 | -29.563
19.0514 | -35.156 | -29.5701
19.0575 | -35.1565 | -29.5774
19.0636 | -35.157 | -29.5844
19.0697 | -35.1575 | -29.5917
19.0758 | -35.158 | -29.5988
19.0819 | -35.1585 | -29.606
19.088 | -35.159 | -29.6132
19.0941 | -35.1595 | -29.6204
19.1002 | -35.16 | -29.6274
19.1063 | -35.1605 | -29.6345
19.1124 | -35.1609 | -29.6418
19.1185 | -35.1614 | -29.649
19.1246 | -35.1619 | -29.6561
19.1308 | -35.1624 | -29.6633
19.1369 | -35.1629 | -29.6705
19.143 | -35.1634 | -29.6776
19.1491 | -35.1639 | -29.6848
19.1552 | -35.1644 | -29.6922
19.1613 | -35.1649 | -29.6993
19.1674 | -35.1654 | -29.7064
19.1735 | -35.1659 | -29.7134
19.1796 | -35.1663 | -29.7204
19.1857 | -35.1668 | -29.7278
19.1918 | -35.1673 | -29.7351
19.1979 | -35.1678 | -29.7425
19.204 | -35.1683 | -29.7494
19.2101 | -35.1688 | -29.7565
19.2162 | -35.1693 | -29.7636
19.2223 | -35.1698 | -29.7708
19.2284 | -35.1703 | -29.778
19.2345 | -35.1708 | -29.785
19.2406 | -35.1713 | -29.7918
19.2467 | -35.1718 | -29.799
19.2528 | -35.1723 | -29.8063
19.2589 | -35.1727 | -29.8135
19.265 | -35.1732 | -29.8206
19.2712 | -35.1737 | -29.8275
19.2773 | -35.1742 | -29.8346
19.2834 | -35.1747 | -29.8417
19.2895 | -35.1752 | -29.849
19.2956 | -35.1757 | -29.8561
19.3017 | -35.1762 | -29.8632
19.3078 | -35.1766 | -29.8704
19.3139 | -35.1771 | -29.8772
19.32 | -35.1776 | -29.8843
19.3261 | -35.1781 | -29.8914
19.3322 | -35.1786 | -29.8985
19.3383 | -35.1791 | -29.9055
19.3444 | -35.1796 | -29.9125
19.3505 | -35.18 | -29.9195
19.3566 | -35.1805 | -29.9266
19.3627 | -35.181 | -29.9336
19.3688 | -35.1815 | -29.9406
19.3749 | -35.182 | -29.9478
19.381 | -35.1824 | -29.9548
19.3871 | -35.1829 | -29.9619
19.3932 | -35.1834 | -29.9689
19.3993 | -35.1839 | -29.9761
19.4054 | -35.1844 | -29.9833
19.4115 | -35.1848 | -29.9904
19.4177 | -35.1853 | -29.9973
19.4238 | -35.1858 | -30.0043
19.4299 | -35.1863 | -30.0113
19.436 | -35.1868 | -30.0185
19.4421 | -35.1873 | -30.0255
19.4482 | -35.1878 | -30.0327
19.4543 | -35.1883 | -30.0394
19.4604 | -35.1888 | -30.0465
19.4665 | -35.1893 | -30.0535
19.4726 | -35.1897 | -30.0604
19.4787 | -35.1902 | -30.0675
19.4848 | -35.1907 | -30.0744
19.4909 | -35.1912 | -30.0813
19.497 | -35.1916 | -30.0881
19.5031 | -35.1921 | -30.0952
19.5092 | -35.1926 | -30.1023
19.5153 | -35.1931 | -30.1092
19.5214 | -35.1936 | -30.1161
19.5275 | -35.194 | -30.1231
19.5336 | -35.1945 | -30.1301
19.5397 | -35.195 | -30.1372
19.5458 | -35.1955 | -30.1443
19.5519 | -35.196 | -30.1512
19.5581 | -35.1965 | -30.1581
19.5642 | -35.1969 | -30.1651
19.5703 | -35.1974 | -30.1722
19.5764 | -35.1979 | -30.1793
19.5825 | -35.1984 | -30.1863
19.5886 | -35.1989 | -30.1934
19.5947 | -35.1994 | -30.2002
19.6008 | -35.1999 | -30.2072
19.6069 | -35.2004 | -30.2144
19.613 | -35.2009 | -30.2214
19.6191 | -35.2013 | -30.2285
19.6252 | -35.2018 | -30.2355
19.6313 | -35.2023 | -30.2422
19.6374 | -35.2028 | -30.2492
19.6435 | -35.2032 | -30.2561
19.6496 | -35.2038 | -30.263
19.6557 | -35.2043 | -30.27
19.6618 | -35.2047 | -30.2768
19.6679 | -35.2052 | -30.2839
19.674 | -35.2057 | -30.2909
19.6801 | -35.2062 | -30.2978
19.6862 | -35.2067 | -30.3047
19.6923 | -35.2072 | -30.3117
19.6984 | -35.2077 | -30.3184
19.7046 | -35.2081 | -30.3254
19.7107 | -35.2086 | -30.3323
19.7168 | -35.2091 | -30.3393
19.7229 | -35.2096 | -30.3462
19.729 | -35.2101 | -30.3531
19.7351 | -35.2106 | -30.3597
19.7412 | -35.2111 | -30.3669
19.7473 | -35.2115 | -30.3739
19.7534 | -35.212 | -30.3809
19.7595 | -35.2125 | -30.388
19.7656 | -35.213 | -30.3947
19.7717 | -35.2135 | -30.4012
19.7778 | -35.2139 | -30.408
19.7839 | -35.2144 | -30.4149
19.79 | -35.2149 | -30.4218
19.7961 | -35.2154 | -30.4288
19.8022 | -35.2158 | -30.4353
19.8083 | -35.2164 | -30.4422
19.8144 | -35.2168 | -30.449
19.8205 | -35.2173 | -30.4559
19.8266 | -35.2178 | -30.4628
19.8327 | -35.2183 | -30.4696
19.8388 | -35.2188 | -30.4761
19.845 | -35.2193 | -30.483
19.8511 | -35.2198 | -30.4899
19.8572 | -35.2203 | -30.4968
19.8633 | -35.2207 | -30.5039
19.8694 | -35.2212 | -30.511
19.8755 | -35.2216 | -30.5178
19.8816 | -35.2221 | -30.5249
19.8877 | -35.2226 | -30.5318
19.8938 | -35.2231 | -30.5388
19.8999 | -35.2235 | -30.5457
19.906 | -35.224 | -30.5526
19.9121 | -35.2245 | -30.5593
19.9182 | -35.225 | -30.5663
19.9243 | -35.2254 | -30.5731
19.9304 | -35.2259 | -30.58
19.9365 | -35.2264 | -30.587
19.9426 | -35.2269 | -30.5936
19.9487 | -35.2274 | -30.6006
19.9548 | -35.2278 | -30.6073
19.9609 | -35.2283 | -30.6142
19.967 | -35.2288 | -30.6213
19.9731 | -35.2293 | -30.6284
19.9792 | -35.2298 | -30.6349
19.9853 | -35.2303 | -30.6417
19.9915 | -35.2307 | -30.6486
19.9976 | -35.2312 | -30.6555
20.0037 | -35.2317 | -30.6625
20.0098 | -35.2321 | -30.6694
20.0159 | -35.2326 | -30.6761
20.022 | -35.2331 | -30.6829
20.0281 | -35.2336 | -30.6899
20.0342 | -35.234 | -30.6968
20.0403 | -35.2345 | -30.7035
20.0464 | -35.235 | -30.7104
20.0525 | -35.2355 | -30.717
20.0586 | -35.236 | -30.724
20.0647 | -35.2365 | -30.7307
20.0708 | -35.2369 | -30.7377
20.0769 | -35.2374 | -30.7446
20.083 | -35.2379 | -30.7511
20.0891 | -35.2384 | -30.7581
20.0952 | -35.2389 | -30.7649
20.1013 | -35.2393 | -30.7719
20.1074 | -35.2398 | -30.7788
20.1135 | -35.2403 | -30.7856
20.1196 | -35.2408 | -30.7921
20.1257 | -35.2412 | -30.799
20.1319 | -35.2417 | -30.8059
20.138 | -35.2422 | -30.8129
20.1441 | -35.2426 | -30.82
20.1502 | -35.2431 | -30.8272
20.1563 | -35.2436 | -30.8336
20.1624 | -35.2441 | -30.8403
20.1685 | -35.2445 | -30.8472
20.1746 | -35.245 | -30.8542
20.1807 | -35.2455 | -30.8609
20.1868 | -35.2459 | -30.8677
20.1929 | -35.2464 | -30.8742
20.199 | -35.2469 | -30.8811
20.2051 | -35.2474 | -30.8878
20.2112 | -35.2478 | -30.8945
20.2173 | -35.2483 | -30.9012
20.2234 | -35.2488 | -30.9078
20.2295 | -35.2493 | -30.9145
20.2356 | -35.2497 | -30.9216
20.2417 | -35.2502 | -30.9286
20.2478 | -35.2507 | -30.9354
20.2539 | -35.2511 | -30.9422
20.26 | -35.2516 | -30.9486
20.2661 | -35.2521 | -30.9552
20.2723 | -35.2526 | -30.962
20.2784 | -35.253 | -30.9687
20.2845 | -35.2535 | -30.9753
20.2906 | -35.254 | -30.982
20.2967 | -35.2544 | -30.9887
20.3028 | -35.2549 | -30.9951
20.3089 | -35.2554 | -31.0021
20.315 | -35.2559 | -31.0088
20.3211 | -35.2564 | -31.0157
20.3272 | -35.2568 | -31.0228
20.3333 | -35.2573 | -31.0292
20.3394 | -35.2578 | -31.0358
20.3455 | -35.2583 | -31.0425
20.3516 | -35.2587 | -31.0493
20.3577 | -35.2592 | -31.0559
20.3638 | -35.2597 | -31.0626
20.3699 | -35.2601 | -31.0687
20.376 | -35.2606 | -31.0756
20.3821 | -35.2611 | -31.0825
20.3882 | -35.2615 | -31.0893
20.3943 | -35.262 | -31.0959
20.4004 | -35.2625 | -31.1022
20.4065 | -35.2629 | -31.109
20.4126 | -35.2634 | -31.1157
20.4188 | -35.2639 | -31.1224
20.4249 | -35.2643 | -31.1292
20.431 | -35.2648 | -31.1359
20.4371 | -35.2653 | -31.1422
20.4432 | -35.2657 | -31.149
20.4493 | -35.2662 | -31.1558
20.4554 | -35.2667 | -31.1625
20.4615 | -35.2671 | -31.1693
20.4676 | -35.2676 | -31.1761
20.4737 | -35.2681 | -31.1828
20.4798 | -35.2686 | -31.1895
20.4859 | -35.2691 | -31.1963
20.492 | -35.2696 | -31.2031
20.4981 | -35.27 | -31.2098
20.5042 | -35.2705 | -31.2166
20.5103 | -35.271 | -31.2228
20.5164 | -35.2714 | -31.2295
20.5225 | -35.2719 | -31.2362
20.5286 | -35.2724 | -31.243
20.5347 | -35.2728 | -31.2495
20.5408 | -35.2733 | -31.2559
20.5469 | -35.2738 | -31.2626
20.553 | -35.2743 | -31.2695
20.5592 | -35.2748 | -31.2761
20.5653 | -35.2753 | -31.2829
20.5714 | -35.2757 | -31.2896
20.5775 | -35.2762 | -31.2962
20.5836 | -35.2767 | -31.3029
20.5897 | -35.2772 | -31.3097
20.5958 | -35.2776 | -31.3164
20.6019 | -35.2781 | -31.3229
20.608 | -35.2785 | -31.3296
20.6141 | -35.279 | -31.336
20.6202 | -35.2795 | -31.3426
20.6263 | -35.28 | -31.3494
20.6324 | -35.2804 | -31.3559
20.6385 | -35.2809 | -31.3626
20.6446 | -35.2814 | -31.3693
20.6507 | -35.2819 | -31.3758
20.6568 | -35.2823 | -31.3826
20.6629 | -35.2828 | -31.3893
20.669 | -35.2833 | -31.3959
20.6751 | -35.2837 | -31.4023
20.6812 | -35.2842 | -31.4088
20.6873 | -35.2846 | -31.4155
20.6934 | -35.2851 | -31.4222
20.6995 | -35.2856 | -31.4286
20.7057 | -35.286 | -31.4355
20.7118 | -35.2865 | -31.4421
20.7179 | -35.287 | -31.4486
20.724 | -35.2874 | -31.4554
20.7301 | -35.2879 | -31.4622
20.7362 | -35.2884 | -31.4688
20.7423 | -35.2889 | -31.4754
20.7484 | -35.2894 | -31.482
20.7545 | -35.2898 | -31.4885
20.7606 | -35.2903 | -31.4953
20.7667 | -35.2908 | -31.5021
20.7728 | -35.2912 | -31.5086
20.7789 | -35.2917 | -31.5153
20.785 | -35.2922 | -31.5219
20.7911 | -35.2927 | -31.5281
20.7972 | -35.2931 | -31.535
20.8033 | -35.2936 | -31.542
20.8094 | -35.2941 | -31.5486
20.8155 | -35.2945 | -31.5553
20.8216 | -35.295 | -31.5619
20.8277 | -35.2954 | -31.5682
20.8338 | -35.2959 | -31.5748
20.8399 | -35.2963 | -31.5817
20.8461 | -35.2968 | -31.5881
20.8522 | -35.2973 | -31.5948
20.8583 | -35.2977 | -31.6014
20.8644 | -35.2982 | -31.6083
20.8705 | -35.2987 | -31.6152
20.8766 | -35.2991 | -31.6218
20.8827 | -35.2996 | -31.6284
20.8888 | -35.3001 | -31.635
20.8949 | -35.3006 | -31.6413
20.901 | -35.3011 | -31.6481
20.9071 | -35.3015 | -31.6543
20.9132 | -35.302 | -31.661
20.9193 | -35.3025 | -31.6675
20.9254 | -35.303 | -31.6743
20.9315 | -35.3034 | -31.6806
20.9376 | -35.3039 | -31.6872
20.9437 | -35.3044 | -31.6939
20.9498 | -35.3049 | -31.7004
20.9559 | -35.3053 | -31.7069
20.962 | -35.3058 | -31.7134
20.9681 | -35.3063 | -31.7199
20.9742 | -35.3068 | -31.7266
20.9803 | -35.3073 | -31.7331
20.9864 | -35.3077 | -31.7399
20.9926 | -35.3082 | -31.7469
20.9987 | -35.3087 | -31.7533
21.0048 | -35.3092 | -31.7597
21.0109 | -35.3096 | -31.7665
21.017 | -35.3101 | -31.773
21.0231 | -35.3106 | -31.7796
21.0292 | -35.311 | -31.7861
21.0353 | -35.3115 | -31.7927
21.0414 | -35.312 | -31.7993
21.0475 | -35.3125 | -31.8061
21.0536 | -35.3129 | -31.8125
21.0597 | -35.3134 | -31.819
21.0658 | -35.3139 | -31.8256
21.0719 | -35.3144 | -31.8322
21.078 | -35.3149 | -31.8386
21.0841 | -35.3153 | -31.8455
21.0902 | -35.3157 | -31.8521
21.0963 | -35.3162 | -31.8581
21.1024 | -35.3169 | -31.8655
21.1085 | -35.3174 | -31.8718
21.1146 | -35.3178 | -31.8783
21.1207 | -35.3183 | -31.8846
21.1268 | -35.3188 | -31.8911
21.133 | -35.3193 | -31.8976
21.1391 | -35.3197 | -31.9041
21.1452 | -35.3202 | -31.9105
21.1513 | -35.3207 | -31.917
21.1574 | -35.3212 | -31.9236
21.1635 | -35.3217 | -31.9301
21.1696 | -35.3221 | -31.9365
21.1757 | -35.3226 | -31.9429
21.1818 | -35.3231 | -31.9496
21.1879 | -35.3235 | -31.9563
21.194 | -35.324 | -31.963
21.2001 | -35.3245 | -31.9695
21.2062 | -35.3249 | -31.9762
21.2123 | -35.3253 | -31.9826
21.2184 | -35.3258 | -31.989
21.2245 | -35.3263 | -31.9955
21.2306 | -35.3268 | -32.0019
21.2367 | -35.3273 | -32.0083
21.2428 | -35.3277 | -32.0149
21.2489 | -35.3282 | -32.0214
21.255 | -35.3286 | -32.0279
21.2611 | -35.3291 | -32.0346
21.2672 | -35.3296 | -32.0413
21.2733 | -35.3301 | -32.0477
21.2795 | -35.3306 | -32.0538
21.2856 | -35.331 | -32.0602
21.2917 | -35.3315 | -32.0665
21.2978 | -35.3319 | -32.0727
21.3039 | -35.3324 | -32.0789
21.31 | -35.3329 | -32.0855
21.3161 | -35.3334 | -32.092
21.3222 | -35.3339 | -32.0989
21.3283 | -35.3344 | -32.1053
21.3344 | -35.3348 | -32.1118
21.3405 | -35.3353 | -32.1184
21.3466 | -35.3358 | -32.1251
21.3527 | -35.3362 | -32.1316
21.3588 | -35.3367 | -32.1383
21.3649 | -35.3372 | -32.1447
21.371 | -35.3377 | -32.1513
21.3771 | -35.3382 | -32.1576
21.3832 | -35.3386 | -32.1641
21.3893 | -35.3391 | -32.1705
21.3954 | -35.3395 | -32.1769
21.4015 | -35.34 | -32.1831
21.4076 | -35.3405 | -32.1897
21.4137 | -35.3409 | -32.196
21.4199 | -35.3414 | -32.2024
21.426 | -35.3418 | -32.2087
21.4321 | -35.3423 | -32.2151
21.4382 | -35.3427 | -32.2215
21.4443 | -35.3432 | -32.2279
21.4504 | -35.3436 | -32.2344
21.4565 | -35.3441 | -32.2407
21.4626 | -35.3445 | -32.2476
21.4687 | -35.345 | -32.2542
21.4748 | -35.3455 | -32.2606
21.4809 | -35.3459 | -32.267
21.487 | -35.3464 | -32.2734
21.4931 | -35.3469 | -32.2795
21.4992 | -35.3473 | -32.2858
21.5053 | -35.3478 | -32.2924
21.5114 | -35.3483 | -32.299
21.5175 | -35.3487 | -32.3053
21.5236 | -35.3492 | -32.312
21.5297 | -35.3496 | -32.3183
21.5358 | -35.3501 | -32.3248
21.5419 | -35.3506 | -32.3314
21.548 | -35.351 | -32.338
21.5541 | -35.3515 | -32.3446
21.5602 | -35.3519 | -32.3512
21.5664 | -35.3524 | -32.3575
21.5725 | -35.3529 | -32.3638
21.5786 | -35.3534 | -32.3703
21.5847 | -35.3538 | -32.3769
21.5908 | -35.3543 | -32.3831
21.5969 | -35.3548 | -32.3892
21.603 | -35.3552 | -32.3958
21.6091 | -35.3557 | -32.4024
21.6152 | -35.3561 | -32.4088
21.6213 | -35.3566 | -32.4152
21.6274 | -35.3571 | -32.4217
21.6335 | -35.3575 | -32.428
21.6396 | -35.358 | -32.4346
21.6457 | -35.3585 | -32.4411
21.6518 | -35.3589 | -32.4475
21.6579 | -35.3594 | -32.4539
21.664 | -35.3599 | -32.4604
21.6701 | -35.3604 | -32.4666
21.6762 | -35.3609 | -32.4729
21.6823 | -35.3614 | -32.4793
21.6884 | -35.3618 | -32.4858
21.6945 | -35.3622 | -32.4924
21.7006 | -35.3627 | -32.4991
21.7068 | -35.3631 | -32.5054
21.7129 | -35.3636 | -32.5119
21.719 | -35.3641 | -32.5184
21.7251 | -35.3645 | -32.5248
21.7312 | -35.3649 | -32.5311
21.7373 | -35.3654 | -32.5372
21.7434 | -35.3659 | -32.5436
21.7495 | -35.3664 | -32.5501
21.7556 | -35.3668 | -32.5567
21.7617 | -35.3673 | -32.5634
21.7678 | -35.3678 | -32.5696
21.7739 | -35.3682 | -32.5758
21.78 | -35.3687 | -32.5825
21.7861 | -35.3691 | -32.5889
21.7922 | -35.3696 | -32.5951
21.7983 | -35.37 | -32.6014
21.8044 | -35.3705 | -32.6079
21.8105 | -35.3709 | -32.6141
21.8166 | -35.3714 | -32.6204
21.8227 | -35.3719 | -32.6268
21.8288 | -35.3723 | -32.6332
21.8349 | -35.3728 | -32.6396
21.841 | -35.3733 | -32.6461
21.8471 | -35.3737 | -32.6523
21.8533 | -35.3742 | -32.6585
21.8594 | -35.3747 | -32.6649
21.8655 | -35.3751 | -32.6713
21.8716 | -35.3756 | -32.6778
21.8777 | -35.3761 | -32.6842
21.8838 | -35.3765 | -32.6906
21.8899 | -35.377 | -32.6971
21.896 | -35.3775 | -32.7033
21.9021 | -35.3779 | -32.7099
21.9082 | -35.3783 | -32.7163
21.9143 | -35.3788 | -32.7225
21.9204 | -35.3793 | -32.7286
21.9265 | -35.3797 | -32.735
21.9326 | -35.3802 | -32.7413
21.9387 | -35.3807 | -32.7478
21.9448 | -35.3811 | -32.7544
21.9509 | -35.3816 | -32.7604
21.957 | -35.382 | -32.7668
21.9631 | -35.3825 | -32.7734
21.9692 | -35.383 | -32.7798
21.9753 | -35.3835 | -32.7863
21.9814 | -35.384 | -32.7927
21.9875 | -35.3844 | -32.7989
21.9937 | -35.3849 | -32.8053
21.9998 | -35.3853 | -32.8118
22.0059 | -35.3857 | -32.8182
22.012 | -35.3862 | -32.8245
22.0181 | -35.3867 | -32.831
22.0242 | -35.3872 | -32.8373
22.0303 | -35.3876 | -32.8435
22.0364 | -35.3881 | -32.8498
22.0425 | -35.3886 | -32.8562
22.0486 | -35.3891 | -32.8625
22.0547 | -35.3895 | -32.869
22.0608 | -35.39 | -32.8753
22.0669 | -35.3905 | -32.8816
22.073 | -35.3909 | -32.8879
22.0791 | -35.3914 | -32.8943
22.0852 | -35.3919 | -32.901
22.0913 | -35.3923 | -32.9073
22.0974 | -35.3928 | -32.9137
22.1035 | -35.3932 | -32.92
22.1096 | -35.3937 | -32.9265
22.1157 | -35.3942 | -32.9329
22.1218 | -35.3946 | -32.9392
22.1279 | -35.3951 | -32.9453
22.134 | -35.3956 | -32.9518
22.1402 | -35.3961 | -32.958
22.1463 | -35.3965 | -32.9642
22.1524 | -35.3969 | -32.9704
22.1585 | -35.3974 | -32.9769
22.1646 | -35.3979 | -32.9831
22.1707 | -35.3983 | -32.9895
22.1768 | -35.3988 | -32.996
22.1829 | -35.3993 | -33.0024
22.189 | -35.3997 | -33.0088
22.1951 | -35.4002 | -33.0149
22.2012 | -35.4006 | -33.0213
22.2073 | -35.4011 | -33.0276
22.2134 | -35.4016 | -33.0338
22.2195 | -35.402 | -33.0402
22.2256 | -35.4025 | -33.0467
22.2317 | -35.4029 | -33.0529
22.2378 | -35.4034 | -33.0595
22.2439 | -35.4038 | -33.0659
22.25 | -35.4042 | -33.0721
22.2561 | -35.4047 | -33.0784
22.2622 | -35.4051 | -33.0847
22.2683 | -35.4056 | -33.0909
22.2744 | -35.4061 | -33.0972
22.2806 | -35.4065 | -33.1038
22.2867 | -35.407 | -33.1102
22.2928 | -35.4075 | -33.1163
22.2989 | -35.4079 | -33.1226
22.305 | -35.4084 | -33.129
22.3111 | -35.4088 | -33.1353
22.3172 | -35.4093 | -33.1414
22.3233 | -35.4098 | -33.1477
22.3294 | -35.4103 | -33.154
22.3355 | -35.4108 | -33.16
22.3416 | -35.4112 | -33.1663
22.3477 | -35.4117 | -33.1727
22.3538 | -35.4122 | -33.1791
22.3599 | -35.4126 | -33.1854
22.366 | -35.4131 | -33.1917
22.3721 | -35.4135 | -33.1979
22.3782 | -35.414 | -33.2041
22.3843 | -35.4145 | -33.2105
22.3904 | -35.415 | -33.2169
22.3965 | -35.4154 | -33.2233
22.4026 | -35.4159 | -33.2297
22.4087 | -35.4163 | -33.2359
22.4148 | -35.4168 | -33.2424
22.4209 | -35.4173 | -33.2488
22.4271 | -35.4177 | -33.2552
22.4332 | -35.4182 | -33.2615
22.4393 | -35.4187 | -33.268
22.4454 | -35.4192 | -33.2739
22.4515 | -35.4196 | -33.2802
22.4576 | -35.4201 | -33.2866
22.4637 | -35.4205 | -33.2929
22.4698 | -35.421 | -33.2993
22.4759 | -35.4215 | -33.3055
22.482 | -35.4219 | -33.312
22.4881 | -35.4224 | -33.3182
22.4942 | -35.4228 | -33.3247
22.5003 | -35.4233 | -33.3311
22.5064 | -35.4238 | -33.3376
22.5125 | -35.4242 | -33.3436
22.5186 | -35.4247 | -33.3498
22.5247 | -35.4251 | -33.3561
22.5308 | -35.4256 | -33.3624
22.5369 | -35.426 | -33.3689
22.543 | -35.4265 | -33.3752
22.5491 | -35.4269 | -33.3815
22.5552 | -35.4274 | -33.3878
22.5613 | -35.4278 | -33.3941
22.5675 | -35.4283 | -33.4005
22.5736 | -35.4287 | -33.407
22.5797 | -35.4292 | -33.4133
22.5858 | -35.4297 | -33.4195
22.5919 | -35.4302 | -33.4257
22.598 | -35.4306 | -33.4321
22.6041 | -35.4311 | -33.4383
22.6102 | -35.4316 | -33.4448
22.6163 | -35.432 | -33.4508
22.6224 | -35.4325 | -33.4571
22.6285 | -35.4329 | -33.4635
22.6346 | -35.4334 | -33.4697
22.6407 | -35.4339 | -33.4762
22.6468 | -35.4343 | -33.4825
22.6529 | -35.4348 | -33.4886
22.659 | -35.4353 | -33.4947
22.6651 | -35.4358 | -33.5011
22.6712 | -35.4362 | -33.5075
22.6773 | -35.4367 | -33.5138
22.6834 | -35.4371 | -33.5202
22.6895 | -35.4376 | -33.5262
22.6956 | -35.438 | -33.5323
22.7017 | -35.4385 | -33.5385
22.7079 | -35.439 | -33.5449
22.714 | -35.4394 | -33.5511
22.7201 | -35.4399 | -33.5574
22.7262 | -35.4403 | -33.5635
22.7323 | -35.4408 | -33.5697
22.7384 | -35.4413 | -33.5761
22.7445 | -35.4417 | -33.5823
22.7506 | -35.4422 | -33.5884
22.7567 | -35.4427 | -33.5947
22.7628 | -35.4431 | -33.6008
22.7689 | -35.4436 | -33.6071
22.775 | -35.444 | -33.6133
22.7811 | -35.4445 | -33.6196
22.7872 | -35.4449 | -33.6257
22.7933 | -35.4454 | -33.6316
22.7994 | -35.4459 | -33.6379
22.8055 | -35.4464 | -33.6441
22.8116 | -35.4468 | -33.6502
22.8177 | -35.4473 | -33.6566
22.8238 | -35.4477 | -33.6628
22.8299 | -35.4482 | -33.6688
22.836 | -35.4487 | -33.6751
22.8421 | -35.4491 | -33.6813
22.8482 | -35.4496 | -33.6877
22.8544 | -35.4501 | -33.694
22.8605 | -35.4505 | -33.7001
22.8666 | -35.451 | -33.7062
22.8727 | -35.4515 | -33.7123
22.8788 | -35.4519 | -33.7183
22.8849 | -35.4524 | -33.7243
22.891 | -35.4528 | -33.7305
22.8971 | -35.4533 | -33.7368
22.9032 | -35.4537 | -33.7428
22.9093 | -35.4542 | -33.7491
22.9154 | -35.4547 | -33.7553
22.9215 | -35.4551 | -33.7614
22.9276 | -35.4556 | -33.7679
22.9337 | -35.456 | -33.7739
22.9398 | -35.4565 | -33.7804
22.9459 | -35.457 | -33.7867
22.952 | -35.4575 | -33.793
22.9581 | -35.4579 | -33.7993
22.9642 | -35.4584 | -33.8055
22.9703 | -35.4589 | -33.8114
22.9764 | -35.4593 | -33.8177
22.9825 | -35.4598 | -33.824
22.9886 | -35.4603 | -33.8302
22.9948 | -35.4607 | -33.8363
23.0009 | -35.4611 | -33.8426
23.007 | -35.4616 | -33.8485
23.0131 | -35.4621 | -33.8548
23.0192 | -35.4625 | -33.861
23.0253 | -35.463 | -33.8674
23.0314 | -35.4634 | -33.8736
23.0375 | -35.4639 | -33.8798
23.0436 | -35.4644 | -33.8859
23.0497 | -35.4648 | -33.8922
23.0558 | -35.4653 | -33.8982
23.0619 | -35.4657 | -33.9042
23.068 | -35.4662 | -33.9104
23.0741 | -35.4667 | -33.9163
23.0802 | -35.4672 | -33.9225
23.0863 | -35.4676 | -33.9287
23.0924 | -35.4681 | -33.9349
23.0985 | -35.4685 | -33.941
23.1046 | -35.469 | -33.9473
23.1107 | -35.4694 | -33.9535
23.1168 | -35.4699 | -33.9597
23.1229 | -35.4704 | -33.9658
23.129 | -35.4708 | -33.9716
23.1351 | -35.4713 | -33.9778
23.1413 | -35.4717 | -33.9839
23.1474 | -35.4722 | -33.9899
23.1535 | -35.4726 | -33.996
23.1596 | -35.4731 | -34.0022
23.1657 | -35.4735 | -34.0083
23.1718 | -35.474 | -34.0144
23.1779 | -35.4744 | -34.0204
23.184 | -35.4749 | -34.0262
23.1901 | -35.4754 | -34.0324
23.1962 | -35.4758 | -34.0386
23.2023 | -35.4762 | -34.0448
23.2084 | -35.4767 | -34.051
23.2145 | -35.4772 | -34.0568
23.2206 | -35.4776 | -34.0629
23.2267 | -35.4781 | -34.0692
23.2328 | -35.4785 | -34.0753
23.2389 | -35.479 | -34.0815
23.245 | -35.4795 | -34.0875
23.2511 | -35.48 | -34.0936
23.2572 | -35.4804 | -34.0996
23.2633 | -35.4809 | -34.1056
23.2694 | -35.4813 | -34.1118
23.2755 | -35.4818 | -34.118
23.2817 | -35.4823 | -34.124
23.2878 | -35.4827 | -34.1301
23.2939 | -35.4832 | -34.1363
23.3 | -35.4837 | -34.1425
23.3061 | -35.4841 | -34.1487
23.3122 | -35.4846 | -34.1549
23.3183 | -35.485 | -34.161
23.3244 | -35.4855 | -34.167
23.3305 | -35.486 | -34.1731
23.3366 | -35.4864 | -34.1793
23.3427 | -35.4869 | -34.1853
23.3488 | -35.4874 | -34.1915
23.3549 | -35.4878 | -34.1976
23.361 | -35.4883 | -34.2035
23.3671 | -35.4887 | -34.2098
23.3732 | -35.4892 | -34.2157
23.3793 | -35.4896 | -34.2218
23.3854 | -35.4901 | -34.228
23.3915 | -35.4906 | -34.2339
23.3976 | -35.491 | -34.24
23.4037 | -35.4915 | -34.2461
23.4098 | -35.4919 | -34.252
23.4159 | -35.4924 | -34.2581
23.422 | -35.4929 | -34.2642
23.4282 | -35.4933 | -34.2702
23.4343 | -35.4938 | -34.2763
23.4404 | -35.4942 | -34.2823
23.4465 | -35.4947 | -34.2885
23.4526 | -35.4951 | -34.2946
23.4587 | -35.4956 | -34.3007
23.4648 | -35.496 | -34.3066
23.4709 | -35.4965 | -34.3126
23.477 | -35.4969 | -34.3188
23.4831 | -35.4974 | -34.3249
23.4892 | -35.4979 | -34.331
23.4953 | -35.4983 | -34.3371
23.5014 | -35.4988 | -34.343
23.5075 | -35.4993 | -34.3491
23.5136 | -35.4997 | -34.3551
23.5197 | -35.5002 | -34.361
23.5258 | -35.5006 | -34.3671
23.5319 | -35.5011 | -34.3729
23.538 | -35.5015 | -34.3789
23.5441 | -35.5019 | -34.385
23.5502 | -35.5024 | -34.3911
23.5563 | -35.5028 | -34.3973
23.5624 | -35.5033 | -34.4034
23.5686 | -35.5037 | -34.4093
23.5747 | -35.5042 | -34.4153
23.5808 | -35.5047 | -34.4214
23.5869 | -35.5051 | -34.4273
23.593 | -35.5056 | -34.4336
23.5991 | -35.506 | -34.4396
23.6052 | -35.5065 | -34.4456
23.6113 | -35.5069 | -34.4517
23.6174 | -35.5074 | -34.4578
23.6235 | -35.5078 | -34.4639
23.6296 | -35.5083 | -34.4701
23.6357 | -35.5088 | -34.4761
23.6418 | -35.5092 | -34.4821
23.6479 | -35.5097 | -34.4882
23.654 | -35.5102 | -34.4944
23.6601 | -35.5106 | -34.5004
23.6662 | -35.5111 | -34.5065
23.6723 | -35.5115 | -34.5123
23.6784 | -35.512 | -34.5182
23.6845 | -35.5124 | -34.5242
23.6906 | -35.5129 | -34.5303
23.6967 | -35.5133 | -34.5363
23.7028 | -35.5138 | -34.5424
23.7089 | -35.5142 | -34.5482
23.7151 | -35.5147 | -34.5542
23.7212 | -35.5151 | -34.5603
23.7273 | -35.5156 | -34.5663
23.7334 | -35.516 | -34.5723
23.7395 | -35.5165 | -34.5784
23.7456 | -35.517 | -34.5839
23.7517 | -35.5174 | -34.59
23.7578 | -35.5179 | -34.596
23.7639 | -35.5183 | -34.602
23.77 | -35.5188 | -34.6081
23.7761 | -35.5192 | -34.6141
23.7822 | -35.5197 | -34.62
23.7883 | -35.5201 | -34.626
23.7944 | -35.5206 | -34.6321
23.8005 | -35.521 | -34.6383
23.8066 | -35.5215 | -34.6444
23.8127 | -35.5219 | -34.6502
23.8188 | -35.5224 | -34.6564
23.8249 | -35.5228 | -34.6624
23.831 | -35.5233 | -34.6683
23.8371 | -35.5237 | -34.6746
23.8432 | -35.5241 | -34.6806
23.8493 | -35.5246 | -34.6864
23.8555 | -35.525 | -34.6923
23.8616 | -35.5255 | -34.6984
23.8677 | -35.5259 | -34.7043
23.8738 | -35.5264 | -34.7102
23.8799 | -35.5269 | -34.7164
23.886 | -35.5273 | -34.7221
23.8921 | -35.5277 | -34.7282
23.8982 | -35.5282 | -34.7342
23.9043 | -35.5287 | -34.7402
23.9104 | -35.5291 | -34.746
23.9165 | -35.5296 | -34.7519
23.9226 | -35.53 | -34.7575
23.9287 | -35.5305 | -34.7635
23.9348 | -35.5309 | -34.7695
23.9409 | -35.5313 | -34.7755
23.947 | -35.5318 | -34.7814
23.9531 | -35.5322 | -34.787
23.9592 | -35.5327 | -34.793
23.9653 | -35.5331 | -34.7991
23.9714 | -35.5336 | -34.8048
23.9775 | -35.534 | -34.8108
23.9836 | -35.5345 | -34.8168
23.9897 | -35.5349 | -34.8226
23.9958 | -35.5354 | -34.8286
24.002 | -35.5359 | -34.8345
24.0081 | -35.5363 | -34.8403
24.0142 | -35.5367 | -34.8462
24.0203 | -35.5372 | -34.8522
24.0264 | -35.5377 | -34.858
24.0325 | -35.5381 | -34.8638
24.0386 | -35.5385 | -34.8699
24.0447 | -35.539 | -34.8757
24.0508 | -35.5394 | -34.8815
24.0569 | -35.5399 | -34.8874
24.063 | -35.5403 | -34.893
24.0691 | -35.5408 | -34.8991
24.0752 | -35.5412 | -34.905
24.0813 | -35.5417 | -34.911
24.0874 | -35.5421 | -34.917
24.0935 | -35.5426 | -34.9228
24.0996 | -35.543 | -34.9287
24.1057 | -35.5435 | -34.9345
24.1118 | -35.5439 | -34.9404
24.1179 | -35.5444 | -34.9465
24.124 | -35.5448 | -34.9525
24.1301 | -35.5453 | -34.9582
24.1362 | -35.5457 | -34.9642
24.1424 | -35.5461 | -34.9703
24.1485 | -35.5466 | -34.9764
24.1546 | -35.5471 | -34.9823
24.1607 | -35.5475 | -34.9883
24.1668 | -35.5479 | -34.9941
24.1729 | -35.5484 | -35.0001
24.179 | -35.5488 | -35.0059
24.1851 | -35.5492 | -35.0118
24.1912 | -35.5497 | -35.0176
24.1973 | -35.5501 | -35.0237
24.2034 | -35.5506 | -35.0294
24.2095 | -35.551 | -35.0355
24.2156 | -35.5515 | -35.0416
24.2217 | -35.5519 | -35.0473
24.2278 | -35.5524 | -35.0531
24.2339 | -35.5528 | -35.059
24.24 | -35.5533 | -35.0646
24.2461 | -35.5538 | -35.0704
24.2522 | -35.5542 | -35.0765
24.2583 | -35.5546 | -35.0824
24.2644 | -35.5551 | -35.0885
24.2705 | -35.5555 | -35.0941
24.2766 | -35.556 | -35.0999
24.2827 | -35.5564 | -35.1057
24.2889 | -35.5569 | -35.1117
24.295 | -35.5573 | -35.1176
24.3011 | -35.5578 | -35.1236
24.3072 | -35.5582 | -35.1292
24.3133 | -35.5586 | -35.135
24.3194 | -35.5591 | -35.141
24.3255 | -35.5595 | -35.147
24.3316 | -35.56 | -35.1528
24.3377 | -35.5604 | -35.1585
24.3438 | -35.5608 | -35.1641
24.3499 | -35.5613 | -35.1699
24.356 | -35.5617 | -35.1756
24.3621 | -35.5621 | -35.1815
24.3682 | -35.5626 | -35.1874
24.3743 | -35.563 | -35.1933
24.3804 | -35.5634 | -35.1989
24.3865 | -35.5638 | -35.2049
24.3926 | -35.5643 | -35.2109
24.3987 | -35.5647 | -35.2168
24.4048 | -35.5652 | -35.2226
24.4109 | -35.5656 | -35.2282
24.417 | -35.566 | -35.234
24.4231 | -35.5665 | -35.2398
24.4293 | -35.567 | -35.2458
24.4354 | -35.5674 | -35.2516
24.4415 | -35.5678 | -35.2574
24.4476 | -35.5683 | -35.2629
24.4537 | -35.5687 | -35.2686
24.4598 | -35.5691 | -35.2745
24.4659 | -35.5696 | -35.2804
24.472 | -35.57 | -35.2863
24.4781 | -35.5705 | -35.292
24.4842 | -35.5709 | -35.2977
24.4903 | -35.5713 | -35.3036
24.4964 | -35.5718 | -35.3094
24.5025 | -35.5722 | -35.3151
24.5086 | -35.5727 | -35.3208
24.5147 | -35.5732 | -35.3265
24.5208 | -35.5736 | -35.3322
24.5269 | -35.5741 | -35.3379
24.533 | -35.5745 | -35.3437
24.5391 | -35.575 | -35.3497
24.5452 | -35.5754 | -35.3555
24.5513 | -35.5758 | -35.3609
24.5574 | -35.5763 | -35.3666
24.5635 | -35.5767 | -35.3723
24.5696 | -35.5771 | -35.3781
24.5758 | -35.5776 | -35.3838
24.5819 | -35.578 | -35.3896
24.588 | -35.5785 | -35.3953
24.5941 | -35.5789 | -35.4011
24.6002 | -35.5794 | -35.4069
24.6063 | -35.5798 | -35.4128
24.6124 | -35.5803 | -35.4186
24.6185 | -35.5807 | -35.4244
24.6246 | -35.5811 | -35.4298
24.6307 | -35.5816 | -35.4355
24.6368 | -35.582 | -35.4413
24.6429 | -35.5825 | -35.4471
24.649 | -35.5829 | -35.4529
24.6551 | -35.5834 | -35.4588
24.6612 | -35.5838 | -35.4641
24.6673 | -35.5842 | -35.4699
24.6734 | -35.5847 | -35.4757
24.6795 | -35.5851 | -35.4816
24.6856 | -35.5855 | -35.4875
24.6917 | -35.586 | -35.4933
24.6978 | -35.5864 | -35.4987
24.7039 | -35.5868 | -35.5043
24.71 | -35.5873 | -35.5103
24.7162 | -35.5877 | -35.5161
24.7223 | -35.5882 | -35.5218
24.7284 | -35.5886 | -35.5273
24.7345 | -35.5891 | -35.533
24.7406 | -35.5895 | -35.5388
24.7467 | -35.5899 | -35.5447
24.7528 | -35.5904 | -35.5506
24.7589 | -35.5908 | -35.5565
24.765 | -35.5912 | -35.562
24.7711 | -35.5917 | -35.5678
24.7772 | -35.5921 | -35.5736
24.7833 | -35.5926 | -35.5793
24.7894 | -35.593 | -35.585
24.7955 | -35.5934 | -35.5908
24.8016 | -35.5939 | -35.596
24.8077 | -35.5943 | -35.6017
24.8138 | -35.5947 | -35.6074
24.8199 | -35.5952 | -35.6131
24.826 | -35.5956 | -35.6188
24.8321 | -35.5961 | -35.6245
24.8382 | -35.5965 | -35.63
24.8443 | -35.5969 | -35.6356
24.8504 | -35.5974 | -35.6413
24.8565 | -35.5978 | -35.6472
24.8627 | -35.5983 | -35.6527
24.8688 | -35.5987 | -35.658
24.8749 | -35.5991 | -35.6638
24.881 | -35.5996 | -35.6696
24.8871 | -35.6 | -35.6754
24.8932 | -35.6004 | -35.6812
24.8993 | -35.6009 | -35.6868
24.9054 | -35.6013 | -35.6921
24.9115 | -35.6018 | -35.6977
24.9176 | -35.6022 | -35.7035
24.9237 | -35.6027 | -35.7091
24.9298 | -35.6031 | -35.7148
24.9359 | -35.6035 | -35.7206
24.942 | -35.6039 | -35.726
24.9481 | -35.6044 | -35.7317
24.9542 | -35.6048 | -35.7374
24.9603 | -35.6052 | -35.7432
24.9664 | -35.6056 | -35.7489
24.9725 | -35.606 | -35.7546
24.9786 | -35.6065 | -35.76
24.9847 | -35.6069 | -35.7657
24.9908 | -35.6073 | -35.7715
24.9969 | -35.6077 | -35.777
25.0031 | -35.6081 | -35.7828
25.0092 | -35.6086 | -35.7884
25.0153 | -35.609 | -35.7941
25.0214 | -35.6095 | -35.7999
25.0275 | -35.6099 | -35.8055
25.0336 | -35.6103 | -35.8111
25.0397 | -35.6107 | -35.817
25.0458 | -35.6111 | -35.8226
25.0519 | -35.6116 | -35.8284
25.058 | -35.6119 | -35.834
25.0641 | -35.6124 | -35.8395
25.0702 | -35.6128 | -35.8452
25.0763 | -35.6132 | -35.8509
25.0824 | -35.6136 | -35.8563
25.0885 | -35.614 | -35.8619
25.0946 | -35.6144 | -35.8676
25.1007 | -35.6149 | -35.8733
25.1068 | -35.6153 | -35.8788
25.1129 | -35.6158 | -35.8847
25.119 | -35.6162 | -35.8901
25.1251 | -35.6166 | -35.8956
25.1312 | -35.6171 | -35.9012
25.1373 | -35.6175 | -35.9069
25.1435 | -35.6179 | -35.9127
25.1496 | -35.6184 | -35.9181
25.1557 | -35.6188 | -35.9238
25.1618 | -35.6192 | -35.9294
25.1679 | -35.6197 | -35.9352
25.174 | -35.6201 | -35.9409
25.1801 | -35.6205 | -35.9466
25.1862 | -35.6209 | -35.9521
25.1923 | -35.6213 | -35.9579
25.1984 | -35.6218 | -35.9636
25.2045 | -35.6222 | -35.9692
25.2106 | -35.6226 | -35.9748
25.2167 | -35.623 | -35.9805
25.2228 | -35.6234 | -35.986
25.2289 | -35.6239 | -35.9917
25.235 | -35.6243 | -35.9974
25.2411 | -35.6247 | -36.003
25.2472 | -35.6252 | -36.0086
25.2533 | -35.6256 | -36.0142
25.2594 | -35.626 | -36.0196
25.2655 | -35.6264 | -36.0252
25.2716 | -35.6268 | -36.0308
25.2777 | -35.6273 | -36.0364
25.2838 | -35.6277 | -36.0421
25.29 | -35.6281 | -36.0478
25.2961 | -35.6286 | -36.0531
25.3022 | -35.629 | -36.0586
25.3083 | -35.6294 | -36.0642
25.3144 | -35.6298 | -36.0698
25.3205 | -35.6303 | -36.0754
25.3266 | -35.6307 | -36.0809
25.3327 | -35.6311 | -36.0865
25.3388 | -35.6315 | -36.0921
25.3449 | -35.632 | -36.0977
25.351 | -35.6324 | -36.1033
25.3571 | -35.6328 | -36.1089
25.3632 | -35.6332 | -36.1142
25.3693 | -35.6336 | -36.1199
25.3754 | -35.6341 | -36.1254
25.3815 | -35.6345 | -36.131
25.3876 | -35.6349 | -36.1369
25.3937 | -35.6353 | -36.1424
25.3998 | -35.6357 | -36.1478
25.4059 | -35.6361 | -36.1535
25.412 | -35.6366 | -36.159
25.4181 | -35.637 | -36.1645
25.4242 | -35.6374 | -36.17
25.4304 | -35.6378 | -36.1756
25.4365 | -35.6382 | -36.181
25.4426 | -35.6386 | -36.1867
25.4487 | -35.639 | -36.1923
25.4548 | -35.6394 | -36.198
25.4609 | -35.6399 | -36.2036
25.467 | -35.6403 | -36.209
25.4731 | -35.6407 | -36.2146
25.4792 | -35.6411 | -36.2202
25.4853 | -35.6415 | -36.2258
25.4914 | -35.642 | -36.2315
25.4975 | -35.6424 | -36.237
25.5036 | -35.6428 | -36.2423
25.5097 | -35.6432 | -36.2478
25.5158 | -35.6437 | -36.2535
25.5219 | -35.6441 | -36.2593
25.528 | -35.6445 | -36.2648
25.5341 | -35.645 | -36.2703
25.5402 | -35.6454 | -36.2757
25.5463 | -35.6458 | -36.2813
25.5524 | -35.6462 | -36.2868
25.5585 | -35.6466 | -36.2924
25.5646 | -35.647 | -36.2979
25.5707 | -35.6474 | -36.3035
25.5769 | -35.6479 | -36.3089
25.583 | -35.6483 | -36.3146
25.5891 | -35.6487 | -36.3201
25.5952 | -35.6491 | -36.3258
25.6013 | -35.6495 | -36.3314
25.6074 | -35.6499 | -36.3366
25.6135 | -35.6504 | -36.3421
25.6196 | -35.6508 | -36.3478
25.6257 | -35.6512 | -36.3533
25.6318 | -35.6516 | -36.3588
25.6379 | -35.652 | -36.3644
25.644 | -35.6524 | -36.3696
25.6501 | -35.6528 | -36.3752
25.6562 | -35.6532 | -36.3808
25.6623 | -35.6537 | -36.3863
25.6684 | -35.6541 | -36.3919
25.6745 | -35.6545 | -36.3974
25.6806 | -35.6549 | -36.4027
25.6867 | -35.6553 | -36.4083
25.6928 | -35.6557 | -36.4138
25.6989 | -35.6561 | -36.4194
25.705 | -35.6565 | -36.4251
25.7111 | -35.6569 | -36.4306
25.7173 | -35.6573 | -36.436
25.7234 | -35.6577 | -36.4417
25.7295 | -35.6581 | -36.4472
25.7356 | -35.6585 | -36.4529
25.7417 | -35.6589 | -36.4584
25.7478 | -35.6593 | -36.4637
25.7539 | -35.6597 | -36.4692
25.76 | -35.6602 | -36.4747
25.7661 | -35.6606 | -36.4802
25.7722 | -35.661 | -36.4857
25.7783 | -35.6614 | -36.4913
25.7844 | -35.6618 | -36.4968
25.7905 | -35.6622 | -36.5024
25.7966 | -35.6626 | -36.508
25.8027 | -35.663 | -36.5134
25.8088 | -35.6634 | -36.519
25.8149 | -35.6638 | -36.5246
25.821 | -35.6642 | -36.5299
25.8271 | -35.6646 | -36.5354
25.8332 | -35.6651 | -36.5411
25.8393 | -35.6655 | -36.5466
25.8454 | -35.6659 | -36.5522
25.8515 | -35.6663 | -36.5576
25.8576 | -35.6667 | -36.563
25.8638 | -35.6671 | -36.5686
25.8699 | -35.6675 | -36.574
25.876 | -35.6679 | -36.5796
25.8821 | -35.6683 | -36.5851
25.8882 | -35.6687 | -36.5903
25.8943 | -35.6691 | -36.5959
25.9004 | -35.6695 | -36.6014
25.9065 | -35.6699 | -36.6069
25.9126 | -35.6703 | -36.6124
25.9187 | -35.6707 | -36.6179
25.9248 | -35.6711 | -36.623
25.9309 | -35.6715 | -36.6285
25.937 | -35.6719 | -36.634
25.9431 | -35.6724 | -36.6395
25.9492 | -35.6728 | -36.6449
25.9553 | -35.6731 | -36.6508
25.9614 | -35.6735 | -36.6561
25.9675 | -35.6739 | -36.6616
25.9736 | -35.6743 | -36.6672
25.9797 | -35.6747 | -36.6727
25.9858 | -35.6751 | -36.678
25.9919 | -35.6755 | -36.6837
25.998 | -35.6759 | -36.6888
26.0042 | -35.6763 | -36.6943
26.0103 | -35.6767 | -36.6998
26.0164 | -35.677 | -36.7052
26.0225 | -35.6775 | -36.7106
26.0286 | -35.6778 | -36.7162
26.0347 | -35.6782 | -36.7214
26.0408 | -35.6786 | -36.7269
26.0469 | -35.679 | -36.7324
26.053 | -35.6795 | -36.7378
26.0591 | -35.6799 | -36.7434
26.0652 | -35.6802 | -36.7486
26.0713 | -35.6806 | -36.7539
26.0774 | -35.681 | -36.7593
26.0835 | -35.6814 | -36.7648
26.0896 | -35.6819 | -36.7702
26.0957 | -35.6823 | -36.7758
26.1018 | -35.6827 | -36.7811
26.1079 | -35.683 | -36.7864
26.114 | -35.6834 | -36.792
26.1201 | -35.6838 | -36.7975
26.1262 | -35.6842 | -36.803
26.1323 | -35.6846 | -36.8084
26.1384 | -35.685 | -36.8138
26.1445 | -35.6854 | -36.8193
26.1507 | -35.6858 | -36.8249
26.1568 | -35.6862 | -36.8303
26.1629 | -35.6866 | -36.8358
26.169 | -35.687 | -36.8413
26.1751 | -35.6874 | -36.8464
26.1812 | -35.6878 | -36.852
26.1873 | -35.6882 | -36.8575
26.1934 | -35.6886 | -36.8629
26.1995 | -35.689 | -36.8685
26.2056 | -35.6894 | -36.8737
26.2117 | -35.6898 | -36.879
26.2178 | -35.6902 | -36.8844
26.2239 | -35.6906 | -36.8899
26.23 | -35.691 | -36.8953
26.2361 | -35.6914 | -36.9009
26.2422 | -35.6918 | -36.9062
26.2483 | -35.6921 | -36.9116
26.2544 | -35.6925 | -36.917
26.2605 | -35.6929 | -36.9225
26.2666 | -35.6933 | -36.928
26.2727 | -35.6937 | -36.9334
26.2788 | -35.6941 | -36.9387
26.2849 | -35.6945 | -36.9442
26.2911 | -35.6949 | -36.9496
26.2972 | -35.6953 | -36.9553
26.3033 | -35.6957 | -36.9607
26.3094 | -35.6961 | -36.9663
26.3155 | -35.6965 | -36.9716
26.3216 | -35.6969 | -36.9771
26.3277 | -35.6972 | -36.9827
26.3338 | -35.6976 | -36.9881
26.3399 | -35.698 | -36.9935
26.346 | -35.6984 | -36.9988
26.3521 | -35.6988 | -37.0044
26.3582 | -35.6992 | -37.0097
26.3643 | -35.6995 | -37.0151
26.3704 | -35.6999 | -37.0205
26.3765 | -35.7003 | -37.026
26.3826 | -35.7007 | -37.0312
26.3887 | -35.7011 | -37.0366
26.3948 | -35.7015 | -37.042
26.4009 | -35.7018 | -37.0474
26.407 | -35.7022 | -37.0529
26.4131 | -35.7027 | -37.0585
26.4192 | -35.703 | -37.0635
26.4253 | -35.7034 | -37.069
26.4314 | -35.7038 | -37.0743
26.4376 | -35.7042 | -37.0797
26.4437 | -35.7046 | -37.085
26.4498 | -35.705 | -37.0906
26.4559 | -35.7054 | -37.0959
26.462 | -35.7057 | -37.1015
26.4681 | -35.7061 | -37.1068
26.4742 | -35.7065 | -37.1123
26.4803 | -35.7069 | -37.1178
26.4864 | -35.7073 | -37.123
26.4925 | -35.7077 | -37.1283
26.4986 | -35.7081 | -37.1337
26.5047 | -35.7085 | -37.1391
26.5108 | -35.7089 | -37.1445
26.5169 | -35.7092 | -37.1499
26.523 | -35.7096 | -37.1551
26.5291 | -35.71 | -37.1604
26.5352 | -35.7104 | -37.1658
26.5413 | -35.7108 | -37.1712
26.5474 | -35.7112 | -37.1765
26.5535 | -35.7115 | -37.182
26.5596 | -35.7119 | -37.1873
26.5657 | -35.7123 | -37.1927
26.5718 | -35.7127 | -37.198
26.578 | -35.7131 | -37.2035
26.5841 | -35.7135 | -37.209
26.5902 | -35.7139 | -37.2143
26.5963 | -35.7142 | -37.2195
26.6024 | -35.7146 | -37.225
26.6085 | -35.715 | -37.2304
26.6146 | -35.7154 | -37.2358
26.6207 | -35.7158 | -37.2412
26.6268 | -35.7162 | -37.2466
26.6329 | -35.7165 | -37.2518
26.639 | -35.7169 | -37.2572
26.6451 | -35.7173 | -37.2626
26.6512 | -35.7177 | -37.2678
26.6573 | -35.7181 | -37.2732
26.6634 | -35.7184 | -37.2784
26.6695 | -35.7188 | -37.2839
26.6756 | -35.7191 | -37.2893
26.6817 | -35.7195 | -37.2947
26.6878 | -35.7199 | -37.3002
26.6939 | -35.7203 | -37.3055
26.7 | -35.7207 | -37.3108
26.7061 | -35.721 | -37.316
26.7122 | -35.7214 | -37.3212
26.7183 | -35.7218 | -37.3265
26.7245 | -35.7222 | -37.332
26.7306 | -35.7226 | -37.3375
26.7367 | -35.7229 | -37.3427
26.7428 | -35.7233 | -37.3481
26.7489 | -35.7237 | -37.3534
26.755 | -35.7241 | -37.3587
26.7611 | -35.7244 | -37.3641
26.7672 | -35.7248 | -37.3695
26.7733 | -35.7252 | -37.3748
26.7794 | -35.7256 | -37.3802
26.7855 | -35.7259 | -37.3854
26.7916 | -35.7263 | -37.3908
26.7977 | -35.7267 | -37.3962
26.8038 | -35.7271 | -37.4013
26.8099 | -35.7274 | -37.4066
26.816 | -35.7278 | -37.4119
26.8221 | -35.7282 | -37.4172
26.8282 | -35.7285 | -37.4225
26.8343 | -35.7289 | -37.4278
26.8404 | -35.7293 | -37.433
26.8465 | -35.7297 | -37.4384
26.8526 | -35.7301 | -37.4437
26.8587 | -35.7304 | -37.4491
26.8649 | -35.7308 | -37.4544
26.871 | -35.7312 | -37.4598
26.8771 | -35.7316 | -37.4649
26.8832 | -35.7319 | -37.4702
26.8893 | -35.7323 | -37.4755
26.8954 | -35.7327 | -37.4808
26.9015 | -35.733 | -37.4863
26.9076 | -35.7334 | -37.4915
26.9137 | -35.7338 | -37.4968
26.9198 | -35.7341 | -37.5022
26.9259 | -35.7345 | -37.5076
26.932 | -35.7349 | -37.5128
26.9381 | -35.7353 | -37.5182
26.9442 | -35.7356 | -37.5235
26.9503 | -35.736 | -37.5287
26.9564 | -35.7363 | -37.5339
26.9625 | -35.7367 | -37.5393
26.9686 | -35.7371 | -37.5446
26.9747 | -35.7375 | -37.5498
26.9808 | -35.7379 | -37.555
26.9869 | -35.7382 | -37.5602
26.993 | -35.7386 | -37.5654
26.9991 | -35.739 | -37.5707
27.0052 | -35.7393 | -37.5762
27.0114 | -35.7397 | -37.5816
27.0175 | -35.7401 | -37.5868
27.0236 | -35.7404 | -37.5922
27.0297 | -35.7408 | -37.5975
27.0358 | -35.7412 | -37.6027
27.0419 | -35.7415 | -37.6081
27.048 | -35.7419 | -37.6133
27.0541 | -35.7422 | -37.6185
27.0602 | -35.7426 | -37.6238
27.0663 | -35.7429 | -37.6291
27.0724 | -35.7433 | -37.6345
27.0785 | -35.7437 | -37.6399
27.0846 | -35.744 | -37.645
27.0907 | -35.7444 | -37.6504
27.0968 | -35.7447 | -37.6557
27.1029 | -35.7451 | -37.661
27.109 | -35.7455 | -37.6663
27.1151 | -35.7459 | -37.6715
27.1212 | -35.7462 | -37.6768
27.1273 | -35.7466 | -37.6822
27.1334 | -35.7469 | -37.6876
27.1395 | -35.7473 | -37.693
27.1456 | -35.7477 | -37.6984
27.1518 | -35.748 | -37.7037
27.1579 | -35.7484 | -37.7088
27.164 | -35.7488 | -37.7141
27.1701 | -35.7491 | -37.7195
27.1762 | -35.7495 | -37.7247
27.1823 | -35.7498 | -37.7299
27.1884 | -35.7502 | -37.7351
27.1945 | -35.7505 | -37.7403
27.2006 | -35.7509 | -37.7456
27.2067 | -35.7512 | -37.7511
27.2128 | -35.7516 | -37.7564
27.2189 | -35.752 | -37.7618
27.225 | -35.7523 | -37.7672
27.2311 | -35.7527 | -37.7724
27.2372 | -35.7531 | -37.7777
27.2433 | -35.7534 | -37.7828
27.2494 | -35.7538 | -37.7883
27.2555 | -35.7542 | -37.7936
27.2616 | -35.7545 | -37.7987
27.2677 | -35.7549 | -37.804
27.2738 | -35.7552 | -37.8092
27.2799 | -35.7556 | -37.8146
27.286 | -35.7559 | -37.8198
27.2921 | -35.7563 | -37.8251
27.2983 | -35.7567 | -37.8301
27.3044 | -35.757 | -37.8356
27.3105 | -35.7573 | -37.8408
27.3166 | -35.7577 | -37.8461
27.3227 | -35.7581 | -37.8514
27.3288 | -35.7584 | -37.8566
27.3349 | -35.7587 | -37.8617
27.341 | -35.7591 | -37.867
27.3471 | -35.7594 | -37.8723
27.3532 | -35.7598 | -37.8774
27.3593 | -35.7602 | -37.8827
27.3654 | -35.7605 | -37.888
27.3715 | -35.7608 | -37.8931
27.3776 | -35.7612 | -37.8985
27.3837 | -35.7615 | -37.9036
27.3898 | -35.7619 | -37.9088
27.3959 | -35.7622 | -37.9141
27.402 | -35.7626 | -37.9191
27.4081 | -35.7629 | -37.9243
27.4142 | -35.7633 | -37.9297
27.4203 | -35.7637 | -37.9349
27.4264 | -35.764 | -37.9402
27.4325 | -35.7644 | -37.9455
27.4387 | -35.7647 | -37.9507
27.4448 | -35.765 | -37.956
27.4509 | -35.7654 | -37.9613
27.457 | -35.7658 | -37.9666
27.4631 | -35.7661 | -37.9719
27.4692 | -35.7665 | -37.9771
27.4753 | -35.7668 | -37.9822
27.4814 | -35.7671 | -37.9875
27.4875 | -35.7675 | -37.9927
27.4936 | -35.7678 | -37.9981
27.4997 | -35.7682 | -38.0035
27.5058 | -35.7685 | -38.0088
27.5119 | -35.7689 | -38.014
27.518 | -35.7692 | -38.0193
27.5241 | -35.7696 | -38.0247
27.5302 | -35.7699 | -38.03
27.5363 | -35.7703 | -38.0353
27.5424 | -35.7706 | -38.0403
27.5485 | -35.771 | -38.0455
27.5546 | -35.7714 | -38.0508
27.5607 | -35.7717 | -38.0562
27.5668 | -35.7721 | -38.0614
27.5729 | -35.7724 | -38.0666
27.5791 | -35.7727 | -38.0717
27.5852 | -35.7731 | -38.0768
27.5913 | -35.7734 | -38.0821
27.5974 | -35.7738 | -38.0873
27.6035 | -35.7742 | -38.0927
27.6096 | -35.7745 | -38.098
27.6157 | -35.7748 | -38.1031
27.6218 | -35.7752 | -38.1085
27.6279 | -35.7755 | -38.1138
27.634 | -35.7759 | -38.1191
27.6401 | -35.7762 | -38.1243
27.6462 | -35.7766 | -38.1297
27.6523 | -35.7769 | -38.1347
27.6584 | -35.7772 | -38.1399
27.6645 | -35.7776 | -38.1452
27.6706 | -35.7779 | -38.1504
27.6767 | -35.7783 | -38.1556
27.6828 | -35.7786 | -38.1608
27.6889 | -35.7789 | -38.1659
27.695 | -35.7793 | -38.1711
27.7011 | -35.7796 | -38.1765
27.7072 | -35.78 | -38.1818
27.7133 | -35.7803 | -38.1871
27.7194 | -35.7807 | -38.1921
27.7256 | -35.781 | -38.1975
27.7317 | -35.7813 | -38.2028
27.7378 | -35.7817 | -38.208
27.7439 | -35.782 | -38.2134
27.75 | -35.7824 | -38.2187
27.7561 | -35.7827 | -38.2237
27.7622 | -35.7831 | -38.2289
27.7683 | -35.7834 | -38.2341
27.7744 | -35.7837 | -38.2394
27.7805 | -35.7841 | -38.2446
27.7866 | -35.7844 | -38.2498
27.7927 | -35.7848 | -38.2549
27.7988 | -35.7851 | -38.2603
27.8049 | -35.7855 | -38.2653
27.811 | -35.7858 | -38.2705
27.8171 | -35.7862 | -38.2757
27.8232 | -35.7865 | -38.2808
27.8293 | -35.7869 | -38.2859
27.8354 | -35.7872 | -38.291
27.8415 | -35.7876 | -38.2963
27.8476 | -35.7879 | -38.3015
27.8537 | -35.7882 | -38.3066
27.8598 | -35.7886 | -38.3117
27.866 | -35.7889 | -38.3168
27.8721 | -35.7892 | -38.3222
27.8782 | -35.7896 | -38.3274
27.8843 | -35.7899 | -38.3326
27.8904 | -35.7903 | -38.3378
27.8965 | -35.7906 | -38.343
27.9026 | -35.7909 | -38.3483
27.9087 | -35.7913 | -38.3537
27.9148 | -35.7916 | -38.3589
27.9209 | -35.7919 | -38.3642
27.927 | -35.7922 | -38.3694
27.9331 | -35.7926 | -38.3745
27.9392 | -35.7929 | -38.3796
27.9453 | -35.7933 | -38.3849
27.9514 | -35.7936 | -38.3902
27.9575 | -35.7939 | -38.3955
27.9636 | -35.7943 | -38.4008
27.9697 | -35.7946 | -38.406
27.9758 | -35.7949 | -38.4112
27.9819 | -35.7952 | -38.4164
27.988 | -35.7955 | -38.4216
27.9941 | -35.7959 | -38.4267
28.0002 | -35.7962 | -38.4318
28.0063 | -35.7965 | -38.4371
28.0125 | -35.7969 | -38.4423
28.0186 | -35.7972 | -38.4475
28.0247 | -35.7976 | -38.4527
28.0308 | -35.7979 | -38.4578
28.0369 | -35.7982 | -38.463
28.043 | -35.7986 | -38.4681
28.0491 | -35.7989 | -38.473
28.0552 | -35.7992 | -38.4783
28.0613 | -35.7996 | -38.4833
28.0674 | -35.7999 | -38.4886
28.0735 | -35.8003 | -38.4936

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4745 | -43.5897
21.4908 | -35.5181 | -35.5137
16.4948 | -34.8895 | -28.0786
12.6602 | -34.5242 | -21.7721
9.7171 | -34.3016 | -16.8417
7.4582 | -34.1449 | -12.847
5.7244 | -34.033 | -9.6151
4.3936 | -33.9564 | -7.0134
3.3722 | -33.9398 | -4.956
2.5883 | -33.9317 | -3.3969
1.9866 | -33.9217 | -2.1498
1.5248 | -33.9234 | -1.1531
1.1703 | -33.9281 | -0.50036
1 | -33.924 | -0.17746
0.89822 | -33.9195 | 0.095654
0.68941 | -33.9056 | 0.67651
0.52919 | -33.8941 | 1.1298
0.40612 | -33.8898 | 1.7416
0.31171 | -33.872 | 2.2039
0.23928 | -33.8774 | 2.935
0.18358 | -33.8888 | 3.3542
0.14091 | -33.8732 | 4.2596
0.10819 | -33.8953 | 5.4199
0.083017 | -33.9115 | 6.9081
0.063705 | -33.9679 | 8.7789
0.048923 | -34.0408 | 11.2783
0.037479 | -34.1418 | 14.4757
0.028801 | -34.3256 | 18.4899
0.022078 | -34.6542 | 23.8487
0.017023 | -35.1251 | 30.2607
0.013018 | -35.944 | 37.9729
0.010014 | -37.1177 | 45.8789

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003735 | -5.0175
0.00747 | -25.9143
0.011205 | -33.4683
0.01494 | -38.5906
0.018675 | -41.6045
0.02241 | -43.5536
0.026145 | -44.3555
0.02988 | -45.3823
0.033615 | -47.0323
0.03735 | -48.4807
0.041085 | -50.8595
0.04482 | -52.1553
0.048555 | -52.5024
0.05229 | -53.049
0.056025 | -54.5893
0.05976 | -56.2347
0.063495 | -57.543
0.06723 | -58.3679
0.070965 | -57.8691
0.0747 | -59.3232
0.078435 | -59.3662
0.08217 | -60.2492
0.085905 | -60.9233
0.08964 | -62.7444
0.093375 | -63.3418
0.09711 | -62.5705
0.10085 | -62.4378
0.10458 | -63.5594
0.10831 | -63.7688
0.11205 | -64.1992
0.11578 | -63.9098
0.11952 | -64.7823
0.12326 | -65.2694
0.12699 | -63.8016
0.13073 | -64.2799
0.13446 | -65.1023
0.1382 | -66.8664
0.14193 | -67.8419
0.14566 | -67.1222
0.1494 | -67.6254
0.15313 | -68.3092
0.15687 | -67.5836
0.1606 | -67.7078
0.16434 | -67.6257
0.16808 | -67.264
0.17181 | -68.8959
0.17555 | -69.9206
0.17928 | -70.0383
0.18302 | -70.1148
0.18675 | -70.3272
0.19048 | -70.0698
0.19422 | -71.1862
0.19795 | -70.5811
0.20169 | -70.5551
0.20542 | -70.0716
0.20916 | -69.4433
0.2129 | -70.8048
0.21663 | -71.2874
0.22037 | -72.4049
0.2241 | -72.7442
0.22784 | -73.2792
0.23157 | -73.5562
0.2353 | -72.0206
0.23904 | -72.344
0.24277 | -73.4162
0.24651 | -73.5324
0.25024 | -73.3602
0.25398 | -73.6881
0.25772 | -73.2961
0.26145 | -73.3531
0.26519 | -73.2684
0.26892 | -73.5131
0.27265 | -75.718
0.27639 | -75.3352
0.28013 | -74.7388
0.28386 | -74.6236
0.28759 | -74.8563
0.29133 | -74.2272
0.29507 | -74.7613
0.2988 | -75.1076
0.30253 | -75.9562
0.30627 | -75.119
0.31 | -75.7529
0.31374 | -76.7029
0.31748 | -75.6864
0.32121 | -74.6212
0.32494 | -74.7372
0.32868 | -74.4256
0.33242 | -74.9115
0.33615 | -76.1117
0.33988 | -75.9618
0.34362 | -76.4209
0.34736 | -76.9501
0.35109 | -77.1516
0.35483 | -77.4967
0.35856 | -77.462
0.36229 | -76.9266
0.36603 | -76.0762
0.36977 | -76.4931
0.3735 | -77.6415
0.37723 | -78.1926
0.38097 | -77.1112
0.38471 | -76.4786
0.38844 | -76.302
0.39217 | -76.9796
0.39591 | -77.4357
0.39964 | -78.1688
0.40338 | -77.3816
0.40712 | -77.5604
0.41085 | -78.4481
0.41458 | -78.116
0.41832 | -77.5057
0.42206 | -77.4532
0.42579 | -77.8118
0.42952 | -78.5561
0.43326 | -78.0967
0.437 | -77.6973
0.44073 | -77.1394
0.44446 | -77.5165
0.4482 | -77.5839
0.45193 | -79.0179
0.45567 | -79.2996
0.45941 | -78.4339
0.46314 | -78.8824
0.46687 | -79.8583
0.47061 | -80.3746
0.47435 | -80.8958
0.47808 | -80.6037
0.48181 | -80.4709
0.48555 | -79.7128
0.48929 | -80.123
0.49302 | -81.4213
0.49676 | -81.2456
0.50049 | -78.9407
0.50423 | -78.5933
0.50796 | -79.8714
0.5117 | -79.4665
0.51543 | -79.5621
0.51916 | -79.9237
0.5229 | -80.2426
0.52663 | -80.971
0.53037 | -81.3785
0.53411 | -80.1663
0.53784 | -80.6651
0.54158 | -80.8357
0.54531 | -80.5928
0.54905 | -80.5106
0.55278 | -80.7305
0.55651 | -81.2496
0.56025 | -81.8492
0.56398 | -82.0295
0.56772 | -81.7993
0.57146 | -81.4238
0.57519 | -81.0385
0.57893 | -80.8404
0.58266 | -80.8077
0.58639 | -81.067
0.59013 | -81.5393
0.59386 | -82.0979
0.5976 | -82.6391
0.60133 | -82.5554
0.60507 | -82.2149
0.60881 | -82.2883
0.61254 | -82.3745
0.61628 | -82.4047
0.62001 | -82.7303
0.62374 | -82.9195
0.62748 | -82.8076
0.63121 | -82.5591
0.63495 | -82.429
0.63868 | -82.383
0.64242 | -82.2556
0.64616 | -82.3732
0.64989 | -82.8645
0.65363 | -83.3752
0.65736 | -83.4492
0.66109 | -83.108
0.66483 | -83.0621
0.66856 | -83.4217
0.6723 | -83.5151
0.67604 | -83.3573
0.67977 | -83.204
0.68351 | -83.1605
0.68724 | -83.1809
0.69098 | -83.3957
0.69471 | -82.9743
0.69844 | -82.5095
0.70218 | -82.5719
0.70591 | -82.6778
0.70965 | -82.7359
0.71339 | -83.1136
0.71712 | -83.5465
0.72086 | -83.8643
0.72459 | -84.1752
0.72833 | -84.2779
0.73206 | -84.3889
0.73579 | -84.3709
0.73953 | -83.6317
0.74326 | -83.1347
0.747 | -83.2452
0.75074 | -83.5669
0.75447 | -83.8632
0.75821 | -84.1794
0.76194 | -84.5833
0.76567 | -84.7236
0.76941 | -84.4586
0.77314 | -84.2034
0.77688 | -84.3569
0.78061 | -84.4284
0.78435 | -83.7892
0.78809 | -83.6203
0.79182 | -84.258
0.79556 | -84.7793
0.79929 | -85.0042
0.80302 | -84.927
0.80676 | -84.9549
0.81049 | -85.2842
0.81423 | -85.2042
0.81797 | -84.5809
0.8217 | -84.4742
0.82544 | -85.0011
0.82917 | -85.4227
0.83291 | -85.1918
0.83664 | -84.9664
0.84037 | -85.0728
0.84411 | -84.8746
0.84784 | -84.6129
0.85158 | -84.6315
0.85532 | -84.9441
0.85905 | -85.3578
0.86279 | -85.6401
0.86652 | -85.4463
0.87026 | -85.1843
0.87399 | -85.4789
0.87772 | -85.8891
0.88146 | -85.5798
0.88519 | -84.6864
0.88893 | -84.3748
0.89267 | -84.6702
0.8964 | -85.0311
0.90014 | -85.3438
0.90387 | -85.614
0.9076 | -85.6662
0.91134 | -85.7506
0.91507 | -86.275
0.91881 | -86.4898
0.92254 | -85.9771
0.92628 | -85.4833
0.93002 | -85.1293
0.93375 | -84.8843
0.93749 | -85.0716
0.94122 | -85.6623
0.94495 | -86.4321
0.94869 | -86.3415
0.95242 | -85.4767
0.95616 | -85.3122
0.9599 | -85.9076
0.96363 | -86.4644
0.96737 | -86.5805
0.9711 | -86.1727
0.97484 | -85.8297
0.97857 | -85.8925
0.9823 | -85.972
0.98604 | -86.0356
0.98977 | -86.3481
0.99351 | -86.425
0.99725 | -86.4066
1.001 | -86.2038
1.0047 | -85.8008
1.0085 | -85.5032
1.0122 | -85.5595
1.0159 | -86.0746
1.0197 | -86.5493
1.0234 | -86.977
1.0271 | -87.3815
1.0309 | -87.2433
1.0346 | -86.631
1.0383 | -85.9732
1.0421 | -85.9469
1.0458 | -86.2132
1.0495 | -86.1544
1.0533 | -85.7829
1.057 | -85.5744
1.0607 | -85.8993
1.0645 | -86.5463
1.0682 | -87.2911
1.0719 | -88.2398
1.0757 | -88.2769
1.0794 | -87.9558
1.0832 | -87.4552
1.0869 | -87.0824
1.0906 | -86.9112
1.0944 | -86.997
1.0981 | -87.2978
1.1018 | -87.6688
1.1056 | -87.9498
1.1093 | -88.0288
1.113 | -87.9359
1.1168 | -87.9093
1.1205 | -88.1348
1.1242 | -88.3443
1.128 | -88.1444
1.1317 | -87.7316
1.1354 | -87.4762
1.1392 | -87.4151
1.1429 | -87.511
1.1466 | -87.7893
1.1504 | -88.1701
1.1541 | -88.4493
1.1579 | -88.4827
1.1616 | -88.4097
1.1653 | -88.3438
1.1691 | -88.253
1.1728 | -88.2256
1.1765 | -88.2691
1.1803 | -88.2976
1.184 | -88.3233
1.1877 | -88.2046
1.1915 | -87.8585
1.1952 | -87.542
1.1989 | -87.4882
1.2027 | -87.6175
1.2064 | -87.8262
1.2101 | -88.2451
1.2139 | -88.8221
1.2176 | -89.0776
1.2213 | -88.9153
1.2251 | -88.7591
1.2288 | -88.8165
1.2326 | -89.0875
1.2363 | -89.4638
1.24 | -89.6527
1.2438 | -89.4594
1.2475 | -89.0046
1.2512 | -88.6006
1.255 | -88.5335
1.2587 | -88.7776
1.2624 | -89.0066
1.2662 | -89.0497
1.2699 | -89.048
1.2736 | -89.1102
1.2774 | -89.2584
1.2811 | -89.5106
1.2848 | -89.7553
1.2886 | -89.854
1.2923 | -89.8881
1.296 | -89.7826
1.2998 | -89.3996
1.3035 | -89.0633
1.3073 | -89.0775
1.311 | -89.3683
1.3147 | -89.647
1.3185 | -89.7061
1.3222 | -89.6047
1.3259 | -89.4297
1.3297 | -89.1818
1.3334 | -88.9022
1.3371 | -88.6918
1.3409 | -88.693
1.3446 | -89.0109
1.3483 | -89.5305
1.3521 | -89.77
1.3558 | -89.5035
1.3595 | -89.2252
1.3633 | -89.301
1.367 | -89.5651
1.3707 | -89.6524
1.3745 | -89.4681
1.3782 | -89.3319
1.382 | -89.5585
1.3857 | -89.9613
1.3894 | -89.9711
1.3932 | -89.5249
1.3969 | -89.0905
1.4006 | -88.9909
1.4044 | -89.2708
1.4081 | -89.7175
1.4118 | -90.0242
1.4156 | -90.1067
1.4193 | -90.043
1.423 | -89.8978
1.4268 | -89.8145
1.4305 | -89.9257
1.4342 | -90.1918
1.438 | -90.4879
1.4417 | -90.6321
1.4454 | -90.5092
1.4492 | -90.2773
1.4529 | -90.2046
1.4567 | -90.3045
1.4604 | -90.3918
1.4641 | -90.4513
1.4679 | -90.5247
1.4716 | -90.5479
1.4753 | -90.5487
1.4791 | -90.5567
1.4828 | -90.4901
1.4865 | -90.3195
1.4903 | -90.2157
1.494 | -90.3739
1.4977 | -90.6092
1.5015 | -90.5501
1.5052 | -90.4252
1.5089 | -90.7445
1.5127 | -91.4296
1.5164 | -91.8095
1.5201 | -91.4671
1.5239 | -90.8085
1.5276 | -90.303
1.5314 | -90.0184
1.5351 | -89.7649
1.5388 | -89.4638
1.5426 | -89.267
1.5463 | -89.2653
1.55 | -89.4394
1.5538 | -89.7519
1.5575 | -90.1127
1.5612 | -90.442
1.565 | -90.6934
1.5687 | -90.8153
1.5724 | -90.84
1.5762 | -90.8233
1.5799 | -90.8669
1.5836 | -91.0934
1.5874 | -91.4829
1.5911 | -91.8176
1.5948 | -91.8352
1.5986 | -91.5344
1.6023 | -91.1855
1.6061 | -90.9333
1.6098 | -90.7991
1.6135 | -90.9943
1.6173 | -91.2142
1.621 | -91.4522
1.6247 | -91.5758
1.6285 | -91.4986
1.6322 | -91.174
1.6359 | -90.6421
1.6397 | -90.0885
1.6434 | -89.7822
1.6471 | -89.8726
1.6509 | -90.3221
1.6546 | -90.9327
1.6583 | -91.4513
1.6621 | -91.7729
1.6658 | -91.9354
1.6695 | -92.004
1.6733 | -92.05
1.677 | -92.0682
1.6808 | -92.0214
1.6845 | -91.92
1.6882 | -91.8224
1.692 | -91.7406
1.6957 | -91.6721
1.6994 | -91.6137
1.7032 | -91.5923
1.7069 | -91.608
1.7106 | -91.6156
1.7144 | -91.585
1.7181 | -91.5037
1.7218 | -91.3813
1.7256 | -91.2398
1.7293 | -91.0914
1.733 | -90.9611
1.7368 | -90.8675
1.7405 | -90.8668
1.7442 | -91.0252
1.748 | -91.316
1.7517 | -91.5914
1.7555 | -91.6763
1.7592 | -91.5856
1.7629 | -91.4747
1.7667 | -91.4278
1.7704 | -91.4268
1.7741 | -91.4453
1.7779 | -91.4602
1.7816 | -91.4691
1.7853 | -91.5016
1.7891 | -91.5826
1.7928 | -91.6907
1.7965 | -91.7862
1.8003 | -91.8392
1.804 | -91.8298
1.8077 | -91.7717
1.8115 | -91.7395
1.8152 | -91.7989
1.8189 | -91.9371
1.8227 | -92.0674
1.8264 | -92.1305
1.8302 | -92.1533
1.8339 | -92.1525
1.8376 | -92.141
1.8414 | -92.127
1.8451 | -92.145
1.8488 | -92.2072
1.8526 | -92.2825
1.8563 | -92.3096
1.86 | -92.2717
1.8638 | -92.2335
1.8675 | -92.2688
1.8712 | -92.343
1.875 | -92.3958
1.8787 | -92.3778
1.8824 | -92.3289
1.8862 | -92.2759
1.8899 | -92.2361
1.8936 | -92.2331
1.8974 | -92.2717
1.9011 | -92.3269
1.9049 | -92.4085
1.9086 | -92.5374
1.9123 | -92.6788
1.9161 | -92.7182
1.9198 | -92.6202
1.9235 | -92.4761
1.9273 | -92.4016
1.931 | -92.4033
1.9347 | -92.3937
1.9385 | -92.3394
1.9422 | -92.283
1.9459 | -92.2808
1.9497 | -92.344
1.9534 | -92.4631
1.9571 | -92.6109
1.9609 | -92.7095
1.9646 | -92.6944
1.9683 | -92.5584
1.9721 | -92.3717
1.9758 | -92.2688
1.9796 | -92.3151
1.9833 | -92.4626
1.987 | -92.6086
1.9908 | -92.698
1.9945 | -92.7653
1.9982 | -92.8419
2.002 | -92.9294
2.0057 | -93.0261
2.0094 | -93.1315
2.0132 | -93.216
2.0169 | -93.2047
2.0206 | -93.071
2.0244 | -92.8877
2.0281 | -92.7429
2.0318 | -92.6609
2.0356 | -92.643
2.0393 | -92.6664
2.043 | -92.6979
2.0468 | -92.6794
2.0505 | -92.539
2.0543 | -92.2774
2.058 | -91.9777
2.0617 | -91.792
2.0655 | -91.8296
2.0692 | -92.038
2.0729 | -92.2757
2.0767 | -92.4676
2.0804 | -92.6496
2.0841 | -92.8273
2.0879 | -92.8603
2.0916 | -92.6799
2.0953 | -92.4055
2.0991 | -92.2457
2.1028 | -92.2633
2.1065 | -92.3868
2.1103 | -92.5453
2.114 | -92.7508
2.1177 | -92.9912
2.1215 | -93.2106
2.1252 | -93.3319
2.129 | -93.3287
2.1327 | -93.2567
2.1364 | -93.2393
2.1402 | -93.2998
2.1439 | -93.3941
2.1476 | -93.4395
2.1514 | -93.4003
2.1551 | -93.3141
2.1588 | -93.2915
2.1626 | -93.3488
2.1663 | -93.4912
2.17 | -93.6637
2.1738 | -93.7783
2.1775 | -93.778
2.1812 | -93.6693
2.185 | -93.5078
2.1887 | -93.3693
2.1924 | -93.3091
2.1962 | -93.3125
2.1999 | -93.3133
2.2037 | -93.2871
2.2074 | -93.2712
2.2111 | -93.3205
2.2149 | -93.4255
2.2186 | -93.5309
2.2223 | -93.5755
2.2261 | -93.5756
2.2298 | -93.5721
2.2335 | -93.6022
2.2373 | -93.6774
2.241 | -93.7782
2.2447 | -93.8813
2.2485 | -93.9588
2.2522 | -93.9529
2.2559 | -93.8641
2.2597 | -93.7317
2.2634 | -93.6227
2.2671 | -93.5763
2.2709 | -93.5671
2.2746 | -93.5475
2.2784 | -93.4737
2.2821 | -93.3316
2.2858 | -93.147
2.2896 | -92.9922
2.2933 | -92.9327
2.297 | -92.9753
2.3008 | -93.0879
2.3045 | -93.2286
2.3082 | -93.3675
2.312 | -93.5102
2.3157 | -93.6345
2.3194 | -93.733
2.3232 | -93.8103
2.3269 | -93.8717
2.3306 | -93.9124
2.3344 | -93.918
2.3381 | -93.9
2.3418 | -93.8869
2.3456 | -93.8712
2.3493 | -93.8668
2.3531 | -93.8862
2.3568 | -93.9535
2.3605 | -94.0671
2.3643 | -94.1885
2.368 | -94.2647
2.3717 | -94.2647
2.3755 | -94.2271
2.3792 | -94.1913
2.3829 | -94.1638
2.3867 | -94.1498
2.3904 | -94.139
2.3941 | -94.1683
2.3979 | -94.2604
2.4016 | -94.4008
2.4053 | -94.5504
2.4091 | -94.6582
2.4128 | -94.6874
2.4165 | -94.6254
2.4203 | -94.524
2.424 | -94.4362
2.4278 | -94.3993
2.4315 | -94.412
2.4352 | -94.4181
2.439 | -94.4111
2.4427 | -94.3906
2.4464 | -94.3663
2.4502 | -94.3485
2.4539 | -94.3468
2.4576 | -94.3844
2.4614 | -94.4504
2.4651 | -94.5182
2.4688 | -94.549
2.4726 | -94.5253
2.4763 | -94.4616
2.48 | -94.3652
2.4838 | -94.2724
2.4875 | -94.224
2.4912 | -94.2175
2.495 | -94.2215
2.4987 | -94.2143
2.5025 | -94.2061
2.5062 | -94.2267
2.5099 | -94.2958
2.5137 | -94.4092
2.5174 | -94.5177
2.5211 | -94.5762
2.5249 | -94.5426
2.5286 | -94.4081
2.5323 | -94.2133
2.5361 | -94.0392
2.5398 | -93.9492
2.5435 | -93.9933
2.5473 | -94.1657
2.551 | -94.4164
2.5547 | -94.662
2.5585 | -94.8418
2.5622 | -94.9605
2.5659 | -95.0548
2.5697 | -95.1176
2.5734 | -95.1387
2.5772 | -95.1067
2.5809 | -95.0401
2.5846 | -94.9333
2.5884 | -94.7836
2.5921 | -94.5819
2.5958 | -94.353
2.5996 | -94.1351
2.6033 | -93.9617
2.607 | -93.905
2.6108 | -93.9785
2.6145 | -94.1415
2.6182 | -94.3233
2.622 | -94.4521
2.6257 | -94.4885
2.6294 | -94.4397
2.6332 | -94.3663
2.6369 | -94.3374
2.6406 | -94.3741
2.6444 | -94.4551
2.6481 | -94.5399
2.6519 | -94.6369
2.6556 | -94.7633
2.6593 | -94.9061
2.6631 | -95.0375
2.6668 | -95.1009
2.6705 | -95.1062
2.6743 | -95.0904
2.678 | -95.0802
2.6817 | -95.0936
2.6855 | -95.1256
2.6892 | -95.1872
2.6929 | -95.2743
2.6967 | -95.3327
2.7004 | -95.3601
2.7041 | -95.3416
2.7079 | -95.2714
2.7116 | -95.1799
2.7153 | -95.1021
2.7191 | -95.0788
2.7228 | -95.1151
2.7266 | -95.1972
2.7303 | -95.3017
2.734 | -95.3825
2.7378 | -95.4255
2.7415 | -95.4201
2.7452 | -95.372
2.749 | -95.3046
2.7527 | -95.2417
2.7564 | -95.1902
2.7602 | -95.1542
2.7639 | -95.1389
2.7676 | -95.1397
2.7714 | -95.1678
2.7751 | -95.2176
2.7788 | -95.2669
2.7826 | -95.3012
2.7863 | -95.3142
2.79 | -95.3214
2.7938 | -95.3097
2.7975 | -95.2753
2.8013 | -95.218
2.805 | -95.1495
2.8087 | -95.0881
2.8125 | -95.0451
2.8162 | -95.0443
2.8199 | -95.0732
2.8237 | -95.1145
2.8274 | -95.1585
2.8311 | -95.2079
2.8349 | -95.2238
2.8386 | -95.1888
2.8423 | -95.1388
2.8461 | -95.1095
2.8498 | -95.1338
2.8535 | -95.2281
2.8573 | -95.3865
2.861 | -95.5923
2.8647 | -95.8184
2.8685 | -96.0485
2.8722 | -96.2511
2.876 | -96.4023
2.8797 | -96.4714
2.8834 | -96.4519
2.8872 | -96.3418
2.8909 | -96.1688
2.8946 | -95.9821
2.8984 | -95.8224
2.9021 | -95.7005
2.9058 | -95.6107
2.9096 | -95.555
2.9133 | -95.5401
2.917 | -95.5446
2.9208 | -95.5517
2.9245 | -95.5473
2.9282 | -95.5386
2.932 | -95.5413
2.9357 | -95.5419
2.9394 | -95.54
2.9432 | -95.5516
2.9469 | -95.5855
2.9507 | -95.6502
2.9544 | -95.7387
2.9581 | -95.834
2.9619 | -95.913
2.9656 | -95.9873
2.9693 | -96.0567
2.9731 | -96.1225
2.9768 | -96.1853
2.9805 | -96.23
2.9843 | -96.2648
2.988 | -96.2902
2.9917 | -96.3401
2.9955 | -96.4206
2.9992 | -96.5114
3.0029 | -96.5709
3.0067 | -96.5612
3.0104 | -96.4708
3.0141 | -96.3048
3.0179 | -96.1042
3.0216 | -95.9155
3.0254 | -95.7758
3.0291 | -95.7156
3.0328 | -95.751
3.0366 | -95.882
3.0403 | -96.0847
3.044 | -96.3218
3.0478 | -96.5154
3.0515 | -96.5909
3.0552 | -96.5358
3.059 | -96.4052
3.0627 | -96.2648
3.0664 | -96.1558
3.0702 | -96.1031
3.0739 | -96.122
3.0776 | -96.2127
3.0814 | -96.3413
3.0851 | -96.4708
3.0888 | -96.5672
3.0926 | -96.6178
3.0963 | -96.6382
3.1001 | -96.6597
3.1038 | -96.6719
3.1075 | -96.6508
3.1113 | -96.5973
3.115 | -96.5366
3.1187 | -96.4669
3.1225 | -96.3858
3.1262 | -96.2975
3.1299 | -96.2238
3.1337 | -96.161
3.1374 | -96.1123
3.1411 | -96.0561
3.1449 | -95.9949
3.1486 | -95.933
3.1523 | -95.9004
3.1561 | -95.9091
3.1598 | -95.9426
3.1635 | -95.9802
3.1673 | -96.0426
3.171 | -96.145
3.1748 | -96.2678
3.1785 | -96.3895
3.1822 | -96.5042
3.186 | -96.5878
3.1897 | -96.6512
3.1934 | -96.6858
3.1972 | -96.6774
3.2009 | -96.6002
3.2046 | -96.4929
3.2084 | -96.3947
3.2121 | -96.3335
3.2158 | -96.3067
3.2196 | -96.3192
3.2233 | -96.3531
3.227 | -96.3992
3.2308 | -96.4409
3.2345 | -96.487
3.2382 | -96.5387
3.242 | -96.583
3.2457 | -96.6194
3.2495 | -96.6624
3.2532 | -96.7055
3.2569 | -96.7288
3.2607 | -96.7482
3.2644 | -96.7727
3.2681 | -96.7972
3.2719 | -96.8177
3.2756 | -96.8408
3.2793 | -96.8527
3.2831 | -96.8626
3.2868 | -96.8655
3.2905 | -96.8637
3.2943 | -96.864
3.298 | -96.8798
3.3017 | -96.8953
3.3055 | -96.9114
3.3092 | -96.9375
3.3129 | -96.9617
3.3167 | -96.9729
3.3204 | -96.9685
3.3242 | -96.9504
3.3279 | -96.922
3.3316 | -96.9065
3.3354 | -96.9055
3.3391 | -96.9138
3.3428 | -96.9361
3.3466 | -96.9574
3.3503 | -96.9667
3.354 | -96.9548
3.3578 | -96.9386
3.3615 | -96.9392
3.3652 | -96.9678
3.369 | -97.0246
3.3727 | -97.1016
3.3764 | -97.1913
3.3802 | -97.2695
3.3839 | -97.3143
3.3876 | -97.331
3.3914 | -97.3507
3.3951 | -97.3928
3.3989 | -97.4653
3.4026 | -97.5626
3.4063 | -97.6733
3.4101 | -97.772
3.4138 | -97.835
3.4175 | -97.8433
3.4213 | -97.806
3.425 | -97.7643
3.4287 | -97.7293
3.4325 | -97.7203
3.4362 | -97.7405
3.4399 | -97.7912
3.4437 | -97.8528
3.4474 | -97.9038
3.4511 | -97.9378
3.4549 | -97.9413
3.4586 | -97.9292
3.4623 | -97.9102
3.4661 | -97.8871
3.4698 | -97.8759
3.4736 | -97.8796
3.4773 | -97.8985
3.481 | -97.9249
3.4848 | -97.9659
3.4885 | -98.0105
3.4922 | -98.04
3.496 | -98.0413
3.4997 | -98.0024
3.5034 | -97.937
3.5072 | -97.8631
3.5109 | -97.8062
3.5146 | -97.7705
3.5184 | -97.7608
3.5221 | -97.7835
3.5258 | -97.8078
3.5296 | -97.8171
3.5333 | -97.7981
3.537 | -97.753
3.5408 | -97.6958
3.5445 | -97.6313
3.5483 | -97.5795
3.552 | -97.5445
3.5557 | -97.5346
3.5595 | -97.5557
3.5632 | -97.6064
3.5669 | -97.6835
3.5707 | -97.7825
3.5744 | -97.872
3.5781 | -97.9271
3.5819 | -97.96
3.5856 | -97.9691
3.5893 | -97.9778
3.5931 | -97.9953
3.5968 | -98.0284
3.6005 | -98.0848
3.6043 | -98.1431
3.608 | -98.1953
3.6117 | -98.2296
3.6155 | -98.2223
3.6192 | -98.1771
3.623 | -98.0912
3.6267 | -97.9783
3.6304 | -97.8678
3.6342 | -97.7821
3.6379 | -97.7287
3.6416 | -97.729
3.6454 | -97.7687
3.6491 | -97.8423
3.6528 | -97.9388
3.6566 | -98.0211
3.6603 | -98.0909
3.664 | -98.1473
3.6678 | -98.1795
3.6715 | -98.1912
3.6752 | -98.1862
3.679 | -98.1793
3.6827 | -98.1665
3.6864 | -98.1561
3.6902 | -98.1529
3.6939 | -98.1659
3.6977 | -98.206
3.7014 | -98.2851
3.7051 | -98.3916
3.7089 | -98.5094
3.7126 | -98.6185
3.7163 | -98.6843
3.7201 | -98.7076
3.7238 | -98.6927
3.7275 | -98.6395
3.7313 | -98.5712
3.735 | -98.511
3.7387 | -98.4766
3.7425 | -98.47
3.7462 | -98.491
3.7499 | -98.5305
3.7537 | -98.5736
3.7574 | -98.604
3.7611 | -98.6124
3.7649 | -98.5789
3.7686 | -98.4855
3.7724 | -98.4153
3.7761 | -98.3355
3.7798 | -98.2674
3.7836 | -98.2254
3.7873 | -98.2052
3.791 | -98.1948
3.7948 | -98.1886
3.7985 | -98.1752
3.8022 | -98.1499
3.806 | -98.1202
3.8097 | -98.0895
3.8134 | -98.0556
3.8172 | -98.0307
3.8209 | -98.0164
3.8246 | -98.0144
3.8284 | -98.0184
3.8321 | -98.0247
3.8358 | -98.0388
3.8396 | -98.0457
3.8433 | -98.0221
3.8471 | -97.9862
3.8508 | -97.938
3.8545 | -97.8907
3.8583 | -97.8585
3.862 | -97.8514
3.8657 | -97.8796
3.8695 | -97.9358
3.8732 | -98.0204
3.8769 | -98.1303
3.8807 | -98.2241
3.8844 | -98.2925
3.8881 | -98.3234
3.8919 | -98.3215
3.8956 | -98.2871
3.8993 | -98.2297
3.9031 | -98.1606
3.9068 | -98.1085
3.9105 | -98.076
3.9143 | -98.0674
3.918 | -98.087
3.9218 | -98.1333
3.9255 | -98.2146
3.9292 | -98.3101
3.933 | -98.4153
3.9367 | -98.514
3.9404 | -98.5867
3.9442 | -98.6441
3.9479 | -98.6743
3.9516 | -98.6865
3.9554 | -98.6817
3.9591 | -98.6823
3.9628 | -98.686
3.9666 | -98.7022
3.9703 | -98.7224
3.974 | -98.7284
3.9778 | -98.7287
3.9815 | -98.7168
3.9852 | -98.6801
3.989 | -98.6095
3.9927 | -98.5223
3.9965 | -98.4364
4.0002 | -98.3745
4.0039 | -98.3496
4.0077 | -98.3579
4.0114 | -98.4032
4.0151 | -98.4703
4.0189 | -98.5523
4.0226 | -98.6472
4.0263 | -98.7551
4.0301 | -98.861
4.0338 | -98.9559
4.0375 | -99.028
4.0413 | -99.0558
4.045 | -99.0277
4.0487 | -98.9497
4.0525 | -98.8523
4.0562 | -98.7635
4.0599 | -98.6862
4.0637 | -98.6258
4.0674 | -98.6037
4.0712 | -98.6137
4.0749 | -98.6454
4.0786 | -98.7008
4.0824 | -98.764
4.0861 | -98.823
4.0898 | -98.8616
4.0936 | -98.8766
4.0973 | -98.8612
4.101 | -98.8209
4.1048 | -98.7548
4.1085 | -98.6837
4.1122 | -98.6237
4.116 | -98.5714
4.1197 | -98.5383
4.1234 | -98.5376
4.1272 | -98.5655
4.1309 | -98.612
4.1346 | -98.6872
4.1384 | -98.7849
4.1421 | -98.8891
4.1459 | -99.0059
4.1496 | -99.1226
4.1533 | -99.2336
4.1571 | -99.3261
4.1608 | -99.3902
4.1645 | -99.4354
4.1683 | -99.4572
4.172 | -99.4457
4.1757 | -99.419
4.1795 | -99.3764
4.1832 | -99.3157
4.1869 | -99.2537
4.1907 | -99.1865
4.1944 | -99.1178
4.1981 | -99.0659
4.2019 | -99.0369
4.2056 | -99.0253
4.2093 | -99.004
4.2131 | -98.9661
4.2168 | -98.925
4.2206 | -98.8795
4.2243 | -98.8322
4.228 | -98.7822
4.2318 | -98.7541
4.2355 | -98.759
4.2392 | -98.784
4.243 | -98.8434
4.2467 | -98.9365
4.2504 | -99.0487
4.2542 | -99.1699
4.2579 | -99.2769
4.2616 | -99.3419
4.2654 | -99.3548
4.2691 | -99.321
4.2728 | -99.2644
4.2766 | -99.1891
4.2803 | -99.0896
4.284 | -98.9947
4.2878 | -98.9135
4.2915 | -98.8451
4.2953 | -98.7882
4.299 | -98.7547
4.3027 | -98.7593
4.3065 | -98.7956
4.3102 | -98.8968
4.3139 | -98.9741
4.3177 | -99.0613
4.3214 | -99.1411
4.3251 | -99.1961
4.3289 | -99.2203
4.3326 | -99.2121
4.3363 | -99.1724
4.3401 | -99.0999
4.3438 | -99.0229
4.3475 | -98.9548
4.3513 | -98.9053
4.355 | -98.8847
4.3587 | -98.8945
4.3625 | -98.9256
4.3662 | -98.9711
4.37 | -99.0045
4.3737 | -99.0335
4.3774 | -99.0353
4.3812 | -98.9983
4.3849 | -98.9303
4.3886 | -98.8432
4.3924 | -98.7597
4.3961 | -98.6941
4.3998 | -98.6556
4.4036 | -98.6543
4.4073 | -98.6953
4.411 | -98.7724
4.4148 | -98.8779
4.4185 | -99.0079
4.4222 | -99.1495
4.426 | -99.2959
4.4297 | -99.4393
4.4334 | -99.5739
4.4372 | -99.6921
4.4409 | -99.7911
4.4447 | -99.843
4.4484 | -99.8775
4.4521 | -99.8848
4.4559 | -99.8689
4.4596 | -99.843
4.4633 | -99.8102
4.4671 | -99.785
4.4708 | -99.7794
4.4745 | -99.7914
4.4783 | -99.8167
4.482 | -99.8555
4.4857 | -99.8991
4.4895 | -99.9371
4.4932 | -99.9549
4.4969 | -99.949
4.5007 | -99.9234
4.5044 | -99.8787
4.5081 | -99.8177
4.5119 | -99.7679
4.5156 | -99.7337
4.5194 | -99.7118
4.5231 | -99.7189
4.5268 | -99.7476
4.5306 | -99.7754
4.5343 | -99.8066
4.538 | -99.8297
4.5418 | -99.8373
4.5455 | -99.8336
4.5492 | -99.8172
4.553 | -99.7888
4.5567 | -99.7525
4.5604 | -99.7133
4.5642 | -99.6774
4.5679 | -99.6517
4.5716 | -99.636
4.5754 | -99.6288
4.5791 | -99.6369
4.5828 | -99.6503
4.5866 | -99.6713
4.5903 | -99.682
4.5941 | -99.6755
4.5978 | -99.6567
4.6015 | -99.6348
4.6053 | -99.6057
4.609 | -99.5724
4.6127 | -99.5404
4.6165 | -99.5144
4.6202 | -99.4969
4.6239 | -99.4926
4.6277 | -99.5009
4.6314 | -99.5193
4.6351 | -99.5465
4.6389 | -99.5752
4.6426 | -99.6056
4.6463 | -99.6366
4.6501 | -99.661
4.6538 | -99.6841
4.6575 | -99.706
4.6613 | -99.7441
4.665 | -99.7979
4.6688 | -99.8635
4.6725 | -99.9276
4.6762 | -99.99
4.68 | -100.0481
4.6837 | -100.0976
4.6874 | -100.129
4.6912 | -100.1432
4.6949 | -100.135
4.6986 | -100.1054
4.7024 | -100.0546
4.7061 | -99.9887
4.7098 | -99.9164
4.7136 | -99.8519
4.7173 | -99.7969
4.721 | -99.7491
4.7248 | -99.7029
4.7285 | -99.6609
4.7322 | -99.6238
4.736 | -99.5933
4.7397 | -99.5743
4.7435 | -99.578
4.7472 | -99.6063
4.7509 | -99.6579
4.7547 | -99.721
4.7584 | -99.7933
4.7621 | -99.8622
4.7659 | -99.9167
4.7696 | -99.9558
4.7733 | -99.9701
4.7771 | -99.9551
4.7808 | -99.9285
4.7845 | -99.899
4.7883 | -99.8756
4.792 | -99.8616
4.7957 | -99.8707
4.7995 | -99.8986
4.8032 | -99.9221
4.8069 | -99.9469
4.8107 | -99.9756
4.8144 | -99.9992
4.8182 | -100.0091
4.8219 | -100.005
4.8256 | -99.9901
4.8294 | -99.9664
4.8331 | -99.9467
4.8368 | -99.9378
4.8406 | -99.9403
4.8443 | -99.9571
4.848 | -99.9688
4.8518 | -99.9841
4.8555 | -99.9936
4.8592 | -99.9873
4.863 | -99.9803
4.8667 | -99.9677
4.8704 | -99.9402
4.8742 | -99.9088
4.8779 | -99.8808
4.8816 | -99.8464
4.8854 | -99.8093
4.8891 | -99.7745
4.8929 | -99.7457
4.8966 | -99.7176
4.9003 | -99.6971
4.9041 | -99.6942
4.9078 | -99.7008
4.9115 | -99.7166
4.9153 | -99.7415
4.919 | -99.7741
4.9227 | -99.8118
4.9265 | -99.8473
4.9302 | -99.8784
4.9339 | -99.9108
4.9377 | -99.9507
4.9414 | -99.9818
4.9451 | -100.01
4.9489 | -100.0347
4.9526 | -100.0526
4.9563 | -100.0701
4.9601 | -100.0819
4.9638 | -100.0899
4.9676 | -100.098
4.9713 | -100.1106
4.975 | -100.1441
4.9788 | -100.1818
4.9825 | -100.2102
4.9862 | -100.2386
4.99 | -100.2659
4.9937 | -100.2743
4.9974 | -100.2648
5.0012 | -100.2367
5.0049 | -100.1955
5.0086 | -100.1431
5.0124 | -100.0896
5.0161 | -100.0423
5.0198 | -100.0126
5.0236 | -100.0024
5.0273 | -100.0088
5.031 | -100.0325
5.0348 | -100.0654
5.0385 | -100.1015
5.0423 | -100.1311
5.046 | -100.153
5.0497 | -100.1623
5.0535 | -100.1586
5.0572 | -100.1515
5.0609 | -100.15
5.0647 | -100.1579
5.0684 | -100.1733
5.0721 | -100.2011
5.0759 | -100.2359
5.0796 | -100.2784
5.0833 | -100.3223
5.0871 | -100.3801
5.0908 | -100.4499
5.0945 | -100.5193
5.0983 | -100.5796
5.102 | -100.6347
5.1057 | -100.6837
5.1095 | -100.7139
5.1132 | -100.737
5.117 | -100.7649
5.1207 | -100.7931
5.1244 | -100.819
5.1282 | -100.8455
5.1319 | -100.8601
5.1356 | -100.8631
5.1394 | -100.8464
5.1431 | -100.8091
5.1468 | -100.7562
5.1506 | -100.6895
5.1543 | -100.6241
5.158 | -100.5588
5.1618 | -100.4898
5.1655 | -100.4191
5.1692 | -100.3546
5.173 | -100.3068
5.1767 | -100.2823
5.1804 | -100.2788
5.1842 | -100.2963
5.1879 | -100.3467
5.1917 | -100.4157
5.1954 | -100.4955
5.1991 | -100.5807
5.2029 | -100.6597
5.2066 | -100.7202
5.2103 | -100.7636
5.2141 | -100.7857
5.2178 | -100.795
5.2215 | -100.8029
5.2253 | -100.8058
5.229 | -100.8075
5.2327 | -100.8107
5.2365 | -100.8286
5.2402 | -100.8605
5.2439 | -100.8956
5.2477 | -100.9328
5.2514 | -100.9684
5.2551 | -101.0048
5.2589 | -101.0372
5.2626 | -101.0695
5.2664 | -101.0884
5.2701 | -101.102
5.2738 | -101.1039
5.2776 | -101.1016
5.2813 | -101.0898
5.285 | -101.0609
5.2888 | -101.0169
5.2925 | -100.9526
5.2962 | -100.8735
5.3 | -100.7905
5.3037 | -100.7172
5.3074 | -100.6545
5.3112 | -100.6034
5.3149 | -100.5735
5.3186 | -100.5707
5.3224 | -100.5916
5.3261 | -100.6283
5.3298 | -100.6909
5.3336 | -100.758
5.3373 | -100.8371
5.3411 | -100.9179
5.3448 | -100.9879
5.3485 | -101.0523
5.3523 | -101.111
5.356 | -101.1526
5.3597 | -101.187
5.3635 | -101.2179
5.3672 | -101.24
5.3709 | -101.2557
5.3747 | -101.2578
5.3784 | -101.2331
5.3821 | -101.1912
5.3859 | -101.113
5.3896 | -101.0581
5.3933 | -100.9967
5.3971 | -100.9334
5.4008 | -100.8748
5.4045 | -100.827
5.4083 | -100.7995
5.412 | -100.7791
5.4158 | -100.7836
5.4195 | -100.8122
5.4232 | -100.8565
5.427 | -100.9171
5.4307 | -100.9693
5.4344 | -101.0214
5.4382 | -101.0643
5.4419 | -101.0854
5.4456 | -101.0837
5.4494 | -101.0609
5.4531 | -101.0216
5.4568 | -100.9731
5.4606 | -100.9301
5.4643 | -100.8895
5.468 | -100.8576
5.4718 | -100.8455
5.4755 | -100.8403
5.4792 | -100.8463
5.483 | -100.8609
5.4867 | -100.8741
5.4905 | -100.8891
5.4942 | -100.8945
5.4979 | -100.8954
5.5017 | -100.8842
5.5054 | -100.8601
5.5091 | -100.8415
5.5129 | -100.8309
5.5166 | -100.832
5.5203 | -100.8342
5.5241 | -100.8487
5.5278 | -100.8774
5.5315 | -100.9147
5.5353 | -100.9583
5.539 | -101.0025
5.5427 | -101.0392
5.5465 | -101.068
5.5502 | -101.0937
5.5539 | -101.1104
5.5577 | -101.1198
5.5614 | -101.1376
5.5652 | -101.16
5.5689 | -101.1848
5.5726 | -101.2146
5.5764 | -101.25
5.5801 | -101.2904
5.5838 | -101.3347
5.5876 | -101.3759
5.5913 | -101.4029
5.595 | -101.4155
5.5988 | -101.4095
5.6025 | -101.3946
5.6062 | -101.3688
5.61 | -101.3358
5.6137 | -101.3016
5.6174 | -101.269
5.6212 | -101.246
5.6249 | -101.2341
5.6286 | -101.242
5.6324 | -101.2652
5.6361 | -101.3033
5.6399 | -101.3527
5.6436 | -101.412
5.6473 | -101.4715
5.6511 | -101.5283
5.6548 | -101.5752
5.6585 | -101.6143
5.6623 | -101.6211
5.666 | -101.6122
5.6697 | -101.5886
5.6735 | -101.5478
5.6772 | -101.4882
5.6809 | -101.4099
5.6847 | -101.3209
5.6884 | -101.2281
5.6921 | -101.137
5.6959 | -101.04
5.6996 | -100.9408
5.7033 | -100.8459
5.7071 | -100.7601
5.7108 | -100.6868
5.7146 | -100.6243
5.7183 | -100.5785
5.722 | -100.5507
5.7258 | -100.5496
5.7295 | -100.5722
5.7332 | -100.6149
5.737 | -100.6768
5.7407 | -100.7582
5.7444 | -100.8456
5.7482 | -100.9378
5.7519 | -101.0254
5.7556 | -101.099
5.7594 | -101.1535
5.7631 | -101.2014
5.7668 | -101.2305
5.7706 | -101.2395
5.7743 | -101.235
5.778 | -101.222
5.7818 | -101.2106
5.7855 | -101.1992
5.7893 | -101.1929
5.793 | -101.1843
5.7967 | -101.1807
5.8005 | -101.1838
5.8042 | -101.2007
5.8079 | -101.2328
5.8117 | -101.2716
5.8154 | -101.327
5.8191 | -101.3886
5.8229 | -101.4518
5.8266 | -101.5058
5.8303 | -101.559
5.8341 | -101.6028
5.8378 | -101.6326
5.8415 | -101.6566
5.8453 | -101.6805
5.849 | -101.688
5.8527 | -101.6845
5.8565 | -101.6763
5.8602 | -101.6586
5.864 | -101.6373
5.8677 | -101.606
5.8714 | -101.5671
5.8752 | -101.5214
5.8789 | -101.4748
5.8826 | -101.4312
5.8864 | -101.3928
5.8901 | -101.3646
5.8938 | -101.3483
5.8976 | -101.3477
5.9013 | -101.3521
5.905 | -101.3734
5.9088 | -101.4061
5.9125 | -101.4432
5.9162 | -101.4713
5.92 | -101.5013
5.9237 | -101.5146
5.9274 | -101.5297
5.9312 | -101.5415
5.9349 | -101.5488
5.9387 | -101.5555
5.9424 | -101.5658
5.9461 | -101.5706
5.9499 | -101.576
5.9536 | -101.5763
5.9573 | -101.5742
5.9611 | -101.5706
5.9648 | -101.5651
5.9685 | -101.5548
5.9723 | -101.5363
5.976 | -101.5177
5.9797 | -101.492
5.9835 | -101.4673
5.9872 | -101.4551
5.9909 | -101.4508
5.9947 | -101.4522
5.9984 | -101.4568
6.0021 | -101.4716
6.0059 | -101.4897
6.0096 | -101.5036
6.0134 | -101.5126
6.0171 | -101.5217
6.0208 | -101.5302
6.0246 | -101.534
6.0283 | -101.5356
6.032 | -101.5351
6.0358 | -101.5249
6.0395 | -101.5121
6.0432 | -101.4937
6.047 | -101.4802
6.0507 | -101.4671
6.0544 | -101.4583
6.0582 | -101.4589
6.0619 | -101.4729
6.0656 | -101.4903
6.0694 | -101.5155
6.0731 | -101.5526
6.0768 | -101.5914
6.0806 | -101.6304
6.0843 | -101.6803
6.0881 | -101.7261
6.0918 | -101.7658
6.0955 | -101.8002
6.0993 | -101.8343
6.103 | -101.8633
6.1067 | -101.8855
6.1105 | -101.9026
6.1142 | -101.9144
6.1179 | -101.9206
6.1217 | -101.9195
6.1254 | -101.9148
6.1291 | -101.9115
6.1329 | -101.8961
6.1366 | -101.8708
6.1403 | -101.8343
6.1441 | -101.7961
6.1478 | -101.7536
6.1515 | -101.7032
6.1553 | -101.6624
6.159 | -101.6342
6.1628 | -101.6061
6.1665 | -101.5838
6.1702 | -101.5664
6.174 | -101.5446
6.1777 | -101.527
6.1814 | -101.5165
6.1852 | -101.5097
6.1889 | -101.5019
6.1926 | -101.4954
6.1964 | -101.4981
6.2001 | -101.5052
6.2038 | -101.5109
6.2076 | -101.5181
6.2113 | -101.5278
6.215 | -101.5415
6.2188 | -101.5602
6.2225 | -101.5763
6.2262 | -101.603
6.23 | -101.6381
6.2337 | -101.6785
6.2375 | -101.7158
6.2412 | -101.7555
6.2449 | -101.8024
6.2487 | -101.841
6.2524 | -101.8716
6.2561 | -101.8993
6.2599 | -101.9219
6.2636 | -101.9376
6.2673 | -101.9483
6.2711 | -101.9376
6.2748 | -101.9192
6.2785 | -101.8943
6.2823 | -101.8594
6.286 | -101.819
6.2897 | -101.7689
6.2935 | -101.7205
6.2972 | -101.6843
6.3009 | -101.6558
6.3047 | -101.6352
6.3084 | -101.624
6.3122 | -101.6167
6.3159 | -101.6248
6.3196 | -101.637
6.3234 | -101.6512
6.3271 | -101.6684
6.3308 | -101.6874
6.3346 | -101.7064
6.3383 | -101.7311
6.342 | -101.7627
6.3458 | -101.8046
6.3495 | -101.8491
6.3532 | -101.8922
6.357 | -101.941
6.3607 | -101.9893
6.3644 | -102.037
6.3682 | -102.0764
6.3719 | -102.1008
6.3756 | -102.1101
6.3794 | -102.1055
6.3831 | -102.0945
6.3869 | -102.0782
6.3906 | -102.0524
6.3943 | -102.0258
6.3981 | -102.0059
6.4018 | -101.9883
6.4055 | -101.9777
6.4093 | -101.9657
6.413 | -101.9612
6.4167 | -101.9672
6.4205 | -101.985
6.4242 | -102.0036
6.4279 | -102.0205
6.4317 | -102.0243
6.4354 | -102.0217
6.4391 | -102.01
6.4429 | -101.9877
6.4466 | -101.9631
6.4503 | -101.9309
6.4541 | -101.8996
6.4578 | -101.8735
6.4616 | -101.8559
6.4653 | -101.8708
6.469 | -101.8784
6.4728 | -101.891
6.4765 | -101.9065
6.4802 | -101.9216
6.484 | -101.9367
6.4877 | -101.9539
6.4914 | -101.9722
6.4952 | -101.9934
6.4989 | -102.0194
6.5026 | -102.0473
6.5064 | -102.0839
6.5101 | -102.1225
6.5138 | -102.1631
6.5176 | -102.2051
6.5213 | -102.2539
6.525 | -102.3007
6.5288 | -102.3489
6.5325 | -102.395
6.5363 | -102.4369
6.54 | -102.4768
6.5437 | -102.5079
6.5475 | -102.5371
6.5512 | -102.5548
6.5549 | -102.5629
6.5587 | -102.5673
6.5624 | -102.5676
6.5661 | -102.5591
6.5699 | -102.5426
6.5736 | -102.5303
6.5773 | -102.5159
6.5811 | -102.5
6.5848 | -102.4873
6.5885 | -102.4766
6.5923 | -102.4659
6.596 | -102.4591
6.5997 | -102.4587
6.6035 | -102.4619
6.6072 | -102.4668
6.611 | -102.4765
6.6147 | -102.4913
6.6184 | -102.5128
6.6222 | -102.5426
6.6259 | -102.5812
6.6296 | -102.6296
6.6334 | -102.6758
6.6371 | -102.7265
6.6408 | -102.7785
6.6446 | -102.8171
6.6483 | -102.843
6.652 | -102.8622
6.6558 | -102.8707
6.6595 | -102.8675
6.6632 | -102.8464
6.667 | -102.8241
6.6707 | -102.7928
6.6744 | -102.7469
6.6782 | -102.7072
6.6819 | -102.6643
6.6857 | -102.6116
6.6894 | -102.5633
6.6931 | -102.5204
6.6969 | -102.4791
6.7006 | -102.4378
6.7043 | -102.4026
6.7081 | -102.3745
6.7118 | -102.3505
6.7155 | -102.3255
6.7193 | -102.3114
6.723 | -102.305
6.7267 | -102.2995
6.7305 | -102.2932
6.7342 | -102.2992
6.7379 | -102.3081
6.7417 | -102.3175
6.7454 | -102.3348
6.7491 | -102.355
6.7529 | -102.3703
6.7566 | -102.3863
6.7604 | -102.3986
6.7641 | -102.4133
6.7678 | -102.4223
6.7716 | -102.4368
6.7753 | -102.4514
6.779 | -102.4578
6.7828 | -102.4654
6.7865 | -102.4772
6.7902 | -102.4881
6.794 | -102.4922
6.7977 | -102.4979
6.8014 | -102.5047
6.8052 | -102.5048
6.8089 | -102.5056
6.8126 | -102.5082
6.8164 | -102.509
6.8201 | -102.5027
6.8238 | -102.5
6.8276 | -102.5065
6.8313 | -102.5138
6.8351 | -102.5192
6.8388 | -102.5222
6.8425 | -102.5229
6.8463 | -102.5216
6.85 | -102.5149
6.8537 | -102.5095
6.8575 | -102.5047
6.8612 | -102.4993
6.8649 | -102.4901
6.8687 | -102.4841
6.8724 | -102.4753
6.8761 | -102.4705
6.8799 | -102.4695
6.8836 | -102.4739
6.8873 | -102.4723
6.8911 | -102.4723
6.8948 | -102.4769
6.8985 | -102.4753
6.9023 | -102.4792
6.906 | -102.4879
6.9098 | -102.4907
6.9135 | -102.492
6.9172 | -102.4997
6.921 | -102.5037
6.9247 | -102.5029
6.9284 | -102.5047
6.9322 | -102.5094
6.9359 | -102.5105
6.9396 | -102.51
6.9434 | -102.5072
6.9471 | -102.4986
6.9508 | -102.4841
6.9546 | -102.4684
6.9583 | -102.4517
6.962 | -102.4352
6.9658 | -102.4284
6.9695 | -102.418
6.9732 | -102.4126
6.977 | -102.4121
6.9807 | -102.4132
6.9845 | -102.4183
6.9882 | -102.4323
6.9919 | -102.4478
6.9957 | -102.4578
6.9994 | -102.472
7.0031 | -102.4787
7.0069 | -102.4852
7.0106 | -102.4929
7.0143 | -102.4946
7.0181 | -102.4905
7.0218 | -102.4838
7.0255 | -102.4757
7.0293 | -102.4645
7.033 | -102.4459
7.0367 | -102.4296
7.0405 | -102.415
7.0442 | -102.4014
7.0479 | -102.392
7.0517 | -102.3903
7.0554 | -102.3945
7.0592 | -102.3987
7.0629 | -102.4104
7.0666 | -102.4326
7.0704 | -102.4551
7.0741 | -102.4843
7.0778 | -102.5139
7.0816 | -102.5467
7.0853 | -102.5809
7.089 | -102.6146
7.0928 | -102.6446
7.0965 | -102.6698
7.1002 | -102.6988
7.104 | -102.7216
7.1077 | -102.7373
7.1114 | -102.7511
7.1152 | -102.7605
7.1189 | -102.7668
7.1226 | -102.7814
7.1264 | -102.7969
7.1301 | -102.8081
7.1339 | -102.8242
7.1376 | -102.8465
7.1413 | -102.8721
7.1451 | -102.8972
7.1488 | -102.9243
7.1525 | -102.955
7.1563 | -102.9848
7.16 | -103.0137
7.1637 | -103.0425
7.1675 | -103.0668
7.1712 | -103.0848
7.1749 | -103.1003
7.1787 | -103.1165
7.1824 | -103.128
7.1861 | -103.1276
7.1899 | -103.1237
7.1936 | -103.1178
7.1973 | -103.1128
7.2011 | -103.1102
7.2048 | -103.1037
7.2086 | -103.0921
7.2123 | -103.0822
7.216 | -103.0778
7.2198 | -103.0778
7.2235 | -103.0793
7.2272 | -103.0824
7.231 | -103.0847
7.2347 | -103.0889
7.2384 | -103.0987
7.2422 | -103.109
7.2459 | -103.113
7.2496 | -103.1216
7.2534 | -103.1334
7.2571 | -103.1453
7.2608 | -103.1564
7.2646 | -103.1667
7.2683 | -103.1741
7.272 | -103.1768
7.2758 | -103.1809
7.2795 | -103.1832
7.2833 | -103.1793
7.287 | -103.1689
7.2907 | -103.1581
7.2945 | -103.1551
7.2982 | -103.1434
7.3019 | -103.1269
7.3057 | -103.1119
7.3094 | -103.0906
7.3131 | -103.0651
7.3169 | -103.0389
7.3206 | -103.0049
7.3243 | -102.9703
7.3281 | -102.9345
7.3318 | -102.9008
7.3355 | -102.8755
7.3393 | -102.8584
7.343 | -102.8494
7.3467 | -102.8446
7.3505 | -102.8461
7.3542 | -102.86
7.358 | -102.875
7.3617 | -102.8883
7.3654 | -102.902
7.3692 | -102.9183
7.3729 | -102.9313
7.3766 | -102.9457
7.3804 | -102.9602
7.3841 | -102.9763
7.3878 | -102.9917
7.3916 | -103.0031
7.3953 | -103.0168
7.399 | -103.0369
7.4028 | -103.0537
7.4065 | -103.0705
7.4102 | -103.091
7.414 | -103.1056
7.4177 | -103.1282
7.4214 | -103.1518
7.4252 | -103.1712
7.4289 | -103.1899
7.4327 | -103.199
7.4364 | -103.2015
7.4401 | -103.2026
7.4439 | -103.1943
7.4476 | -103.1832
7.4513 | -103.1719
7.4551 | -103.1602
7.4588 | -103.1493
7.4625 | -103.1319
7.4663 | -103.114
7.47 | -103.0957
7.4737 | -103.0775
7.4775 | -103.0571
7.4812 | -103.0428
7.4849 | -103.0287
7.4887 | -103.018
7.4924 | -103.0101
7.4961 | -103.0091
7.4999 | -103.017
7.5036 | -103.0298
7.5074 | -103.0406
7.5111 | -103.0563
7.5148 | -103.0755
7.5186 | -103.0978
7.5223 | -103.1248
7.526 | -103.149
7.5298 | -103.1768
7.5335 | -103.2021
7.5372 | -103.2215
7.541 | -103.2121
7.5447 | -103.2164
7.5484 | -103.2155
7.5522 | -103.2067
7.5559 | -103.1956
7.5596 | -103.1799
7.5634 | -103.1645
7.5671 | -103.1481
7.5708 | -103.1324
7.5746 | -103.1278
7.5783 | -103.1335
7.5821 | -103.1453
7.5858 | -103.1622
7.5895 | -103.1906
7.5933 | -103.22
7.597 | -103.2531
7.6007 | -103.2903
7.6045 | -103.3301
7.6082 | -103.3677
7.6119 | -103.3972
7.6157 | -103.4177
7.6194 | -103.4292
7.6231 | -103.4349
7.6269 | -103.4291
7.6306 | -103.4092
7.6343 | -103.3795
7.6381 | -103.3431
7.6418 | -103.2988
7.6455 | -103.2538
7.6493 | -103.212
7.653 | -103.1698
7.6568 | -103.1356
7.6605 | -103.1042
7.6642 | -103.0824
7.668 | -103.0672
7.6717 | -103.0629
7.6754 | -103.0711
7.6792 | -103.0873
7.6829 | -103.1208
7.6866 | -103.1647
7.6904 | -103.2151
7.6941 | -103.2722
7.6978 | -103.3368
7.7016 | -103.4061
7.7053 | -103.4671
7.709 | -103.529
7.7128 | -103.5878
7.7165 | -103.6355
7.7202 | -103.6732
7.724 | -103.7049
7.7277 | -103.7324
7.7315 | -103.7494
7.7352 | -103.7608
7.7389 | -103.7547
7.7427 | -103.7424
7.7464 | -103.7212
7.7501 | -103.6897
7.7539 | -103.6548
7.7576 | -103.6159
7.7613 | -103.5733
7.7651 | -103.526
7.7688 | -103.4809
7.7725 | -103.4356
7.7763 | -103.3898
7.78 | -103.3424
7.7837 | -103.305
7.7875 | -103.2744
7.7912 | -103.2504
7.7949 | -103.2385
7.7987 | -103.2323
7.8024 | -103.2366
7.8062 | -103.2479
7.8099 | -103.262
7.8136 | -103.2853
7.8174 | -103.3192
7.8211 | -103.3603
7.8248 | -103.4006
7.8286 | -103.4359
7.8323 | -103.4666
7.836 | -103.4871
7.8398 | -103.499
7.8435 | -103.5043
7.8472 | -103.5029
7.851 | -103.4936
7.8547 | -103.4885
7.8584 | -103.4726
7.8622 | -103.4515
7.8659 | -103.4316
7.8696 | -103.4089
7.8734 | -103.3922
7.8771 | -103.3789
7.8809 | -103.3703
7.8846 | -103.368
7.8883 | -103.367
7.8921 | -103.3782
7.8958 | -103.3977
7.8995 | -103.4212
7.9033 | -103.4437
7.907 | -103.4712
7.9107 | -103.5065
7.9145 | -103.5401
7.9182 | -103.5763
7.9219 | -103.6144
7.9257 | -103.65
7.9294 | -103.6805
7.9331 | -103.6986
7.9369 | -103.7159
7.9406 | -103.7344
7.9443 | -103.7473
7.9481 | -103.7543
7.9518 | -103.7616
7.9556 | -103.7662
7.9593 | -103.7666
7.963 | -103.7612
7.9668 | -103.7541
7.9705 | -103.7456
7.9742 | -103.7346
7.978 | -103.7118
7.9817 | -103.6835
7.9854 | -103.652
7.9892 | -103.6127
7.9929 | -103.5697
7.9966 | -103.5323
8.0004 | -103.4977
8.0041 | -103.4642
8.0078 | -103.433
8.0116 | -103.4074
8.0153 | -103.3884
8.019 | -103.3796
8.0228 | -103.3776
8.0265 | -103.377
8.0303 | -103.3847
8.034 | -103.3954
8.0377 | -103.4085
8.0415 | -103.4298
8.0452 | -103.452
8.0489 | -103.4764
8.0527 | -103.5021
8.0564 | -103.5303
8.0601 | -103.5632
8.0639 | -103.59
8.0676 | -103.6085
8.0713 | -103.6228
8.0751 | -103.6339
8.0788 | -103.6442
8.0825 | -103.6422
8.0863 | -103.6417
8.09 | -103.6347
8.0937 | -103.6283
8.0975 | -103.6211
8.1012 | -103.6136
8.105 | -103.6058
8.1087 | -103.5949
8.1124 | -103.5818
8.1162 | -103.5631
8.1199 | -103.5471
8.1236 | -103.5378
8.1274 | -103.5279
8.1311 | -103.5185
8.1348 | -103.5116
8.1386 | -103.5087
8.1423 | -103.5016
8.146 | -103.4937
8.1498 | -103.4949
8.1535 | -103.498
8.1572 | -103.5107
8.161 | -103.5243
8.1647 | -103.5389
8.1684 | -103.558
8.1722 | -103.5754
8.1759 | -103.5953
8.1797 | -103.6157
8.1834 | -103.6441
8.1871 | -103.6713
8.1909 | -103.6978
8.1946 | -103.7209
8.1983 | -103.744
8.2021 | -103.7689
8.2058 | -103.7958
8.2095 | -103.8211
8.2133 | -103.8458
8.217 | -103.8681
8.2207 | -103.8956
8.2245 | -103.9151
8.2282 | -103.9325
8.2319 | -103.9524
8.2357 | -103.9704
8.2394 | -103.9887
8.2431 | -104.0027
8.2469 | -104.0109
8.2506 | -104.0165
8.2544 | -104.0167
8.2581 | -104.0099
8.2618 | -104.0009
8.2656 | -103.9929
8.2693 | -103.9835
8.273 | -103.964
8.2768 | -103.9418
8.2805 | -103.9172
8.2842 | -103.8958
8.288 | -103.883
8.2917 | -103.8672
8.2954 | -103.8511
8.2992 | -103.8385
8.3029 | -103.8314
8.3066 | -103.8215
8.3104 | -103.8156
8.3141 | -103.8119
8.3178 | -103.8109
8.3216 | -103.8107
8.3253 | -103.8113
8.3291 | -103.8158
8.3328 | -103.8232
8.3365 | -103.8345
8.3403 | -103.8516
8.344 | -103.8702
8.3477 | -103.889
8.3515 | -103.9078
8.3552 | -103.9284
8.3589 | -103.9508
8.3627 | -103.9677
8.3664 | -103.9792
8.3701 | -103.9852
8.3739 | -103.9831
8.3776 | -103.98
8.3813 | -103.971
8.3851 | -103.9553
8.3888 | -103.9363
8.3925 | -103.913
8.3963 | -103.887
8.4 | -103.8581
8.4038 | -103.83
8.4075 | -103.8061
8.4112 | -103.7833
8.415 | -103.7623
8.4187 | -103.7439
8.4224 | -103.7249
8.4262 | -103.709
8.4299 | -103.7024
8.4336 | -103.6984
8.4374 | -103.6978
8.4411 | -103.6981
8.4448 | -103.695
8.4486 | -103.6967
8.4523 | -103.6995
8.456 | -103.7021
8.4598 | -103.7004
8.4635 | -103.6963
8.4672 | -103.6861
8.471 | -103.6819
8.4747 | -103.6803
8.4785 | -103.6771
8.4822 | -103.6708
8.4859 | -103.6682
8.4897 | -103.6715
8.4934 | -103.6779
8.4971 | -103.6885
8.5009 | -103.6948
8.5046 | -103.7024
8.5083 | -103.712
8.5121 | -103.7189
8.5158 | -103.7238
8.5195 | -103.7333
8.5233 | -103.747
8.527 | -103.7573
8.5307 | -103.7697
8.5345 | -103.7868
8.5382 | -103.8128
8.5419 | -103.8461
8.5457 | -103.88
8.5494 | -103.9125
8.5532 | -103.9443
8.5569 | -103.9787
8.5606 | -104.0143
8.5644 | -104.0524
8.5681 | -104.0866
8.5718 | -104.114
8.5756 | -104.134
8.5793 | -104.1431
8.583 | -104.1509
8.5868 | -104.1528
8.5905 | -104.1486
8.5942 | -104.1373
8.598 | -104.1267
8.6017 | -104.1135
8.6054 | -104.1054
8.6092 | -104.095
8.6129 | -104.0771
8.6166 | -104.0616
8.6204 | -104.0553
8.6241 | -104.0367
8.6279 | -104.0231
8.6316 | -104.0104
8.6353 | -103.9917
8.6391 | -103.9703
8.6428 | -103.9483
8.6465 | -103.9224
8.6503 | -103.8974
8.654 | -103.8714
8.6577 | -103.8459
8.6615 | -103.8213
8.6652 | -103.7956
8.6689 | -103.7665
8.6727 | -103.741
8.6764 | -103.7234
8.6801 | -103.7101
8.6839 | -103.7015
8.6876 | -103.6941
8.6913 | -103.6887
8.6951 | -103.6905
8.6988 | -103.6968
8.7026 | -103.7071
8.7063 | -103.7245
8.71 | -103.7411
8.7138 | -103.7578
8.7175 | -103.7764
8.7212 | -103.793
8.725 | -103.8047
8.7287 | -103.808
8.7324 | -103.8046
8.7362 | -103.8002
8.7399 | -103.7971
8.7436 | -103.7921
8.7474 | -103.7786
8.7511 | -103.763
8.7548 | -103.752
8.7586 | -103.7351
8.7623 | -103.7176
8.766 | -103.7031
8.7698 | -103.6989
8.7735 | -103.6956
8.7773 | -103.695
8.781 | -103.6932
8.7847 | -103.6985
8.7885 | -103.7097
8.7922 | -103.7239
8.7959 | -103.7437
8.7997 | -103.7683
8.8034 | -103.7957
8.8071 | -103.8216
8.8109 | -103.8488
8.8146 | -103.8767
8.8183 | -103.9065
8.8221 | -103.941
8.8258 | -103.9772
8.8295 | -104.0072
8.8333 | -104.0377
8.837 | -104.0664
8.8407 | -104.0928
8.8445 | -104.1177
8.8482 | -104.1399
8.852 | -104.1615
8.8557 | -104.1816
8.8594 | -104.2063
8.8632 | -104.2278
8.8669 | -104.2397
8.8706 | -104.2496
8.8744 | -104.2547
8.8781 | -104.2575
8.8818 | -104.2571
8.8856 | -104.2531
8.8893 | -104.2466
8.893 | -104.2395
8.8968 | -104.2306
8.9005 | -104.2104
8.9042 | -104.1896
8.908 | -104.1739
8.9117 | -104.162
8.9154 | -104.1525
8.9192 | -104.1476
8.9229 | -104.1461
8.9267 | -104.1473
8.9304 | -104.1509
8.9341 | -104.1536
8.9379 | -104.1584
8.9416 | -104.1639
8.9453 | -104.1722
8.9491 | -104.1809
8.9528 | -104.1857
8.9565 | -104.1895
8.9603 | -104.1951
8.964 | -104.1963
8.9677 | -104.1964
8.9715 | -104.197
8.9752 | -104.1987
8.9789 | -104.2017
8.9827 | -104.2002
8.9864 | -104.1965
8.9901 | -104.1978
8.9939 | -104.2015
8.9976 | -104.2092
9.0014 | -104.2209
9.0051 | -104.2364
9.0088 | -104.256
9.0126 | -104.2781
9.0163 | -104.3046
9.02 | -104.3359
9.0238 | -104.3707
9.0275 | -104.4106
9.0312 | -104.4512
9.035 | -104.4924
9.0387 | -104.5317
9.0424 | -104.5714
9.0462 | -104.607
9.0499 | -104.634
9.0536 | -104.6637
9.0574 | -104.6949
9.0611 | -104.7234
9.0648 | -104.7469
9.0686 | -104.7681
9.0723 | -104.7824
9.0761 | -104.7868
9.0798 | -104.7907
9.0835 | -104.791
9.0873 | -104.7845
9.091 | -104.7741
9.0947 | -104.762
9.0985 | -104.7395
9.1022 | -104.7091
9.1059 | -104.6764
9.1097 | -104.6458
9.1134 | -104.6156
9.1171 | -104.588
9.1209 | -104.5603
9.1246 | -104.5373
9.1283 | -104.5164
9.1321 | -104.4992
9.1358 | -104.4835
9.1395 | -104.4732
9.1433 | -104.4622
9.147 | -104.4511
9.1508 | -104.4409
9.1545 | -104.4314
9.1582 | -104.4295
9.162 | -104.4227
9.1657 | -104.4163
9.1694 | -104.4131
9.1732 | -104.4098
9.1769 | -104.4057
9.1806 | -104.4012
9.1844 | -104.4049
9.1881 | -104.408
9.1918 | -104.409
9.1956 | -104.4142
9.1993 | -104.4173
9.203 | -104.4206
9.2068 | -104.4303
9.2105 | -104.4411
9.2142 | -104.4451
9.218 | -104.4514
9.2217 | -104.4586
9.2255 | -104.4598
9.2292 | -104.4599
9.2329 | -104.463
9.2367 | -104.4598
9.2404 | -104.4492
9.2441 | -104.4334
9.2479 | -104.42
9.2516 | -104.4025
9.2553 | -104.3808
9.2591 | -104.357
9.2628 | -104.3355
9.2665 | -104.314
9.2703 | -104.3013
9.274 | -104.2865
9.2777 | -104.2752
9.2815 | -104.2674
9.2852 | -104.26
9.2889 | -104.2525
9.2927 | -104.2453
9.2964 | -104.2383
9.3002 | -104.2315
9.3039 | -104.2313
9.3076 | -104.2295
9.3114 | -104.2304
9.3151 | -104.2327
9.3188 | -104.2394
9.3226 | -104.2422
9.3263 | -104.2436
9.33 | -104.2493
9.3338 | -104.2592
9.3375 | -104.2772
9.3412 | -104.289
9.345 | -104.2986
9.3487 | -104.3113
9.3524 | -104.3315
9.3562 | -104.3514
9.3599 | -104.365
9.3636 | -104.3816
9.3674 | -104.3988
9.3711 | -104.4117
9.3749 | -104.4243
9.3786 | -104.4328
9.3823 | -104.4376
9.3861 | -104.439
9.3898 | -104.4425
9.3935 | -104.4451
9.3973 | -104.4465
9.401 | -104.4452
9.4047 | -104.4404
9.4085 | -104.4291
9.4122 | -104.4223
9.4159 | -104.4145
9.4197 | -104.4012
9.4234 | -104.386
9.4271 | -104.3712
9.4309 | -104.3605
9.4346 | -104.3529
9.4383 | -104.3472
9.4421 | -104.339
9.4458 | -104.3298
9.4496 | -104.3226
9.4533 | -104.3212
9.457 | -104.3216
9.4608 | -104.3253
9.4645 | -104.3352
9.4682 | -104.3435
9.472 | -104.357
9.4757 | -104.3749
9.4794 | -104.3968
9.4832 | -104.4171
9.4869 | -104.4406
9.4906 | -104.4605
9.4944 | -104.4784
9.4981 | -104.4999
9.5018 | -104.5162
9.5056 | -104.5266
9.5093 | -104.5349
9.513 | -104.5375
9.5168 | -104.5368
9.5205 | -104.5291
9.5243 | -104.5135
9.528 | -104.4902
9.5317 | -104.4695
9.5355 | -104.4466
9.5392 | -104.4216
9.5429 | -104.3978
9.5467 | -104.3786
9.5504 | -104.3586
9.5541 | -104.3401
9.5579 | -104.3275
9.5616 | -104.3159
9.5653 | -104.308
9.5691 | -104.303
9.5728 | -104.3027
9.5765 | -104.3074
9.5803 | -104.3145
9.584 | -104.3238
9.5877 | -104.3318
9.5915 | -104.342
9.5952 | -104.3577
9.599 | -104.3708
9.6027 | -104.3839
9.6064 | -104.3984
9.6102 | -104.4066
9.6139 | -104.4148
9.6176 | -104.4281
9.6214 | -104.4379
9.6251 | -104.4415
9.6288 | -104.4463
9.6326 | -104.4519
9.6363 | -104.4584
9.64 | -104.4655
9.6438 | -104.4698
9.6475 | -104.4728
9.6512 | -104.4759
9.655 | -104.4798
9.6587 | -104.4848
9.6624 | -104.4885
9.6662 | -104.4921
9.6699 | -104.4949
9.6737 | -104.4993
9.6774 | -104.5053
9.6811 | -104.5128
9.6849 | -104.5166
9.6886 | -104.5204
9.6923 | -104.5283
9.6961 | -104.5252
9.6998 | -104.526
9.7035 | -104.529
9.7073 | -104.5308
9.711 | -104.5279
9.7147 | -104.5208
9.7185 | -104.5136
9.7222 | -104.5083
9.7259 | -104.504
9.7297 | -104.4987
9.7334 | -104.4881
9.7371 | -104.4749
9.7409 | -104.4664
9.7446 | -104.4574
9.7484 | -104.4461
9.7521 | -104.4356
9.7558 | -104.426
9.7596 | -104.4147
9.7633 | -104.4058
9.767 | -104.4013
9.7708 | -104.3951
9.7745 | -104.3877
9.7782 | -104.3823
9.782 | -104.379
9.7857 | -104.3735
9.7894 | -104.3706
9.7932 | -104.3686
9.7969 | -104.3668
9.8006 | -104.3659
9.8044 | -104.3663
9.8081 | -104.3678
9.8118 | -104.3663
9.8156 | -104.3674
9.8193 | -104.3699
9.8231 | -104.3734
9.8268 | -104.3795
9.8305 | -104.3868
9.8343 | -104.3938
9.838 | -104.4037
9.8417 | -104.4143
9.8455 | -104.4249
9.8492 | -104.4339
9.8529 | -104.4463
9.8567 | -104.463
9.8604 | -104.4828
9.8641 | -104.5014
9.8679 | -104.5185
9.8716 | -104.536
9.8753 | -104.5539
9.8791 | -104.5739
9.8828 | -104.5902
9.8865 | -104.6033
9.8903 | -104.6168
9.894 | -104.6298
9.8978 | -104.6401
9.9015 | -104.6408
9.9052 | -104.6393
9.909 | -104.6368
9.9127 | -104.6328
9.9164 | -104.6258
9.9202 | -104.6177
9.9239 | -104.6077
9.9276 | -104.5969
9.9314 | -104.5867
9.9351 | -104.5778
9.9388 | -104.5715
9.9426 | -104.5661
9.9463 | -104.563
9.95 | -104.5642
9.9538 | -104.5673
9.9575 | -104.5757
9.9612 | -104.5889
9.965 | -104.604
9.9687 | -104.6221
9.9725 | -104.6469
9.9762 | -104.676
9.9799 | -104.7058
9.9837 | -104.7358
9.9874 | -104.7623
9.9911 | -104.7917
9.9949 | -104.8214
9.9986 | -104.8457
10.0023 | -104.8692
10.0061 | -104.8914
10.0098 | -104.9078
10.0135 | -104.9213
10.0173 | -104.9322
10.021 | -104.9396
10.0247 | -104.9406
10.0285 | -104.9391
10.0322 | -104.938
10.0359 | -104.9334
10.0397 | -104.9293
10.0434 | -104.9253
10.0472 | -104.9289
10.0509 | -104.9335
10.0546 | -104.9331
10.0584 | -104.935
10.0621 | -104.942
10.0658 | -104.949
10.0696 | -104.9585
10.0733 | -104.9715
10.077 | -104.989
10.0808 | -105.0012
10.0845 | -105.0108
10.0882 | -105.0216
10.092 | -105.0295
10.0957 | -105.0399
10.0994 | -105.0493
10.1032 | -105.0516
10.1069 | -105.0509
10.1106 | -105.0503
10.1144 | -105.0449
10.1181 | -105.0335
10.1219 | -105.0209
10.1256 | -105.0045
10.1293 | -104.9963
10.1331 | -104.9923
10.1368 | -104.981
10.1405 | -104.9692
10.1443 | -104.9586
10.148 | -104.9502
10.1517 | -104.9482
10.1555 | -104.9507
10.1592 | -104.952
10.1629 | -104.9507
10.1667 | -104.9461
10.1704 | -104.9462
10.1741 | -104.9464
10.1779 | -104.938
10.1816 | -104.9306
10.1853 | -104.9234
10.1891 | -104.9101
10.1928 | -104.8952
10.1966 | -104.8804
10.2003 | -104.8677
10.204 | -104.8501
10.2078 | -104.8341
10.2115 | -104.8212
10.2152 | -104.8115
10.219 | -104.8029
10.2227 | -104.7969
10.2264 | -104.7961
10.2302 | -104.7989
10.2339 | -104.8024
10.2376 | -104.8164
10.2414 | -104.8251
10.2451 | -104.8344
10.2488 | -104.8477
10.2526 | -104.866
10.2563 | -104.8801
10.26 | -104.894
10.2638 | -104.9072
10.2675 | -104.9219
10.2713 | -104.935
10.275 | -104.9493
10.2787 | -104.9543
10.2825 | -104.9548
10.2862 | -104.9564
10.2899 | -104.961
10.2937 | -104.9616
10.2974 | -104.9539
10.3011 | -104.9434
10.3049 | -104.935
10.3086 | -104.9295
10.3123 | -104.9215
10.3161 | -104.9167
10.3198 | -104.9092
10.3235 | -104.8979
10.3273 | -104.8879
10.331 | -104.8792
10.3347 | -104.8659
10.3385 | -104.8582
10.3422 | -104.8528
10.346 | -104.8475
10.3497 | -104.8456
10.3534 | -104.8445
10.3572 | -104.8455
10.3609 | -104.8457
10.3646 | -104.8441
10.3684 | -104.8426
10.3721 | -104.843
10.3758 | -104.8461
10.3796 | -104.8503
10.3833 | -104.8525
10.387 | -104.8534
10.3908 | -104.8483
10.3945 | -104.8415
10.3982 | -104.8342
10.402 | -104.8264
10.4057 | -104.8162
10.4094 | -104.8073
10.4132 | -104.8002
10.4169 | -104.7925
10.4207 | -104.783
10.4244 | -104.7698
10.4281 | -104.7585
10.4319 | -104.7513
10.4356 | -104.743
10.4393 | -104.7315
10.4431 | -104.7181
10.4468 | -104.7089
10.4505 | -104.7043
10.4543 | -104.696
10.458 | -104.6896
10.4617 | -104.6805
10.4655 | -104.6746
10.4692 | -104.6721
10.4729 | -104.6735
10.4767 | -104.6775
10.4804 | -104.6881
10.4841 | -104.6986
10.4879 | -104.7173
10.4916 | -104.7344
10.4954 | -104.7488
10.4991 | -104.7657
10.5028 | -104.7848
10.5066 | -104.8037
10.5103 | -104.8199
10.514 | -104.8367
10.5178 | -104.8548
10.5215 | -104.8699
10.5252 | -104.879
10.529 | -104.8871
10.5327 | -104.8926
10.5364 | -104.8939
10.5402 | -104.8911
10.5439 | -104.8911
10.5476 | -104.8874
10.5514 | -104.8806
10.5551 | -104.8726
10.5588 | -104.8676
10.5626 | -104.8632
10.5663 | -104.8578
10.5701 | -104.8535
10.5738 | -104.8501
10.5775 | -104.8491
10.5813 | -104.8445
10.585 | -104.8385
10.5887 | -104.837
10.5925 | -104.8424
10.5962 | -104.8445
10.5999 | -104.8454
10.6037 | -104.8453
10.6074 | -104.8485
10.6111 | -104.8501
10.6149 | -104.8537
10.6186 | -104.8553
10.6223 | -104.8552
10.6261 | -104.8554
10.6298 | -104.8566
10.6335 | -104.8536
10.6373 | -104.8442
10.641 | -104.8319
10.6448 | -104.8257
10.6485 | -104.8189
10.6522 | -104.8114
10.656 | -104.8024
10.6597 | -104.7933
10.6634 | -104.7856
10.6672 | -104.7779
10.6709 | -104.769
10.6746 | -104.7633
10.6784 | -104.7582
10.6821 | -104.7467
10.6858 | -104.7324
10.6896 | -104.7233
10.6933 | -104.7159
10.697 | -104.705
10.7008 | -104.6915
10.7045 | -104.6756
10.7082 | -104.6616
10.712 | -104.6456
10.7157 | -104.6301
10.7195 | -104.6159
10.7232 | -104.6034
10.7269 | -104.5941
10.7307 | -104.5891
10.7344 | -104.5882
10.7381 | -104.5901
10.7419 | -104.5919
10.7456 | -104.5897
10.7493 | -104.5935
10.7531 | -104.5994
10.7568 | -104.6087
10.7605 | -104.6193
10.7643 | -104.6283
10.768 | -104.6399
10.7717 | -104.6548
10.7755 | -104.6715
10.7792 | -104.6813
10.7829 | -104.6897
10.7867 | -104.7013
10.7904 | -104.7118
10.7942 | -104.7272
10.7979 | -104.7441
10.8016 | -104.7575
10.8054 | -104.7734
10.8091 | -104.7881
10.8128 | -104.8032
10.8166 | -104.8188
10.8203 | -104.833
10.824 | -104.845
10.8278 | -104.8564
10.8315 | -104.8672
10.8352 | -104.8765
10.839 | -104.8864
10.8427 | -104.8978
10.8464 | -104.9052
10.8502 | -104.9075
10.8539 | -104.9047
10.8576 | -104.9042
10.8614 | -104.9049
10.8651 | -104.9011
10.8689 | -104.8993
10.8726 | -104.8961
10.8763 | -104.8923
10.8801 | -104.8914
10.8838 | -104.8919
10.8875 | -104.8952
10.8913 | -104.9004
10.895 | -104.9057
10.8987 | -104.9146
10.9025 | -104.9279
10.9062 | -104.9401
10.9099 | -104.9497
10.9137 | -104.9584
10.9174 | -104.9711
10.9211 | -104.9849
10.9249 | -104.9984
10.9286 | -105.0107
10.9323 | -105.025
10.9361 | -105.0402
10.9398 | -105.0593
10.9436 | -105.0788
10.9473 | -105.0978
10.951 | -105.119
10.9548 | -105.1379
10.9585 | -105.1546
10.9622 | -105.1701
10.966 | -105.1836
10.9697 | -105.202
10.9734 | -105.2216
10.9772 | -105.2386
10.9809 | -105.2517
10.9846 | -105.2615
10.9884 | -105.2715
10.9921 | -105.2796
10.9958 | -105.2855
10.9996 | -105.29
11.0033 | -105.2891
11.007 | -105.2872
11.0108 | -105.2872
11.0145 | -105.2866
11.0183 | -105.2837
11.022 | -105.2778
11.0257 | -105.2737
11.0295 | -105.269
11.0332 | -105.2666
11.0369 | -105.2669
11.0407 | -105.2648
11.0444 | -105.2619
11.0481 | -105.2573
11.0519 | -105.2566
11.0556 | -105.2631
11.0593 | -105.2703
11.0631 | -105.2778
11.0668 | -105.2876
11.0705 | -105.3
11.0743 | -105.311
11.078 | -105.3182
11.0817 | -105.3265
11.0855 | -105.3304
11.0892 | -105.3351
11.093 | -105.339
11.0967 | -105.3391
11.1004 | -105.3321
11.1042 | -105.3212
11.1079 | -105.3091
11.1116 | -105.2951
11.1154 | -105.2802
11.1191 | -105.2662
11.1228 | -105.251
11.1266 | -105.2352
11.1303 | -105.2188
11.134 | -105.2028
11.1378 | -105.1892
11.1415 | -105.1785
11.1452 | -105.1677
11.149 | -105.1544
11.1527 | -105.1435
11.1564 | -105.1369
11.1602 | -105.132
11.1639 | -105.1292
11.1677 | -105.1248
11.1714 | -105.1217
11.1751 | -105.1202
11.1789 | -105.1179
11.1826 | -105.1167
11.1863 | -105.1162
11.1901 | -105.1144
11.1938 | -105.1134
11.1975 | -105.1119
11.2013 | -105.1139
11.205 | -105.1181
11.2087 | -105.1252
11.2125 | -105.1259
11.2162 | -105.1252
11.2199 | -105.125
11.2237 | -105.1244
11.2274 | -105.123
11.2311 | -105.1226
11.2349 | -105.122
11.2386 | -105.1201
11.2424 | -105.1165
11.2461 | -105.1122
11.2498 | -105.1082
11.2536 | -105.1057
11.2573 | -105.105
11.261 | -105.1067
11.2648 | -105.1107
11.2685 | -105.1108
11.2722 | -105.1129
11.276 | -105.117
11.2797 | -105.1231
11.2834 | -105.131
11.2872 | -105.1381
11.2909 | -105.1439
11.2946 | -105.1506
11.2984 | -105.1606
11.3021 | -105.1686
11.3058 | -105.1797
11.3096 | -105.1884
11.3133 | -105.1867
11.3171 | -105.1881
11.3208 | -105.1857
11.3245 | -105.1835
11.3283 | -105.1813
11.332 | -105.1775
11.3357 | -105.1712
11.3395 | -105.1655
11.3432 | -105.1609
11.3469 | -105.1528
11.3507 | -105.1432
11.3544 | -105.1356
11.3581 | -105.1251
11.3619 | -105.1143
11.3656 | -105.1036
11.3693 | -105.0936
11.3731 | -105.0834
11.3768 | -105.0766
11.3805 | -105.0683
11.3843 | -105.0639
11.388 | -105.0628
11.3918 | -105.0645
11.3955 | -105.0664
11.3992 | -105.0692
11.403 | -105.0726
11.4067 | -105.0802
11.4104 | -105.0915
11.4142 | -105.1048
11.4179 | -105.1205
11.4216 | -105.1387
11.4254 | -105.158
11.4291 | -105.1781
11.4328 | -105.2027
11.4366 | -105.2304
11.4403 | -105.2575
11.444 | -105.2823
11.4478 | -105.3117
11.4515 | -105.3427
11.4552 | -105.3717
11.459 | -105.4022
11.4627 | -105.4283
11.4665 | -105.4569
11.4702 | -105.4854
11.4739 | -105.5145
11.4777 | -105.546
11.4814 | -105.5746
11.4851 | -105.5995
11.4889 | -105.6222
11.4926 | -105.6403
11.4963 | -105.6554
11.5001 | -105.6701
11.5038 | -105.6815
11.5075 | -105.687
11.5113 | -105.689
11.515 | -105.6902
11.5187 | -105.6886
11.5225 | -105.6829
11.5262 | -105.6743
11.5299 | -105.662
11.5337 | -105.6474
11.5374 | -105.633
11.5412 | -105.618
11.5449 | -105.6042
11.5486 | -105.5901
11.5524 | -105.5726
11.5561 | -105.5567
11.5598 | -105.5443
11.5636 | -105.5352
11.5673 | -105.5302
11.571 | -105.5312
11.5748 | -105.5324
11.5785 | -105.5334
11.5822 | -105.5359
11.586 | -105.5432
11.5897 | -105.5539
11.5934 | -105.5696
11.5972 | -105.588
11.6009 | -105.6055
11.6046 | -105.623
11.6084 | -105.6383
11.6121 | -105.6546
11.6159 | -105.6698
11.6196 | -105.6779
11.6233 | -105.6855
11.6271 | -105.6903
11.6308 | -105.6952
11.6345 | -105.6926
11.6383 | -105.6842
11.642 | -105.6759
11.6457 | -105.6616
11.6495 | -105.6448
11.6532 | -105.6306
11.6569 | -105.6154
11.6607 | -105.5949
11.6644 | -105.5729
11.6681 | -105.5489
11.6719 | -105.5251
11.6756 | -105.4968
11.6793 | -105.4708
11.6831 | -105.4472
11.6868 | -105.424
11.6906 | -105.4026
11.6943 | -105.3862
11.698 | -105.3729
11.7018 | -105.3579
11.7055 | -105.3469
11.7092 | -105.341
11.713 | -105.3399
11.7167 | -105.3411
11.7204 | -105.3429
11.7242 | -105.3466
11.7279 | -105.3515
11.7316 | -105.355
11.7354 | -105.3536
11.7391 | -105.3538
11.7428 | -105.355
11.7466 | -105.3528
11.7503 | -105.3452
11.754 | -105.3345
11.7578 | -105.3237
11.7615 | -105.3135
11.7653 | -105.3004
11.769 | -105.2914
11.7727 | -105.2812
11.7765 | -105.2695
11.7802 | -105.2611
11.7839 | -105.2499
11.7877 | -105.243
11.7914 | -105.2365
11.7951 | -105.2278
11.7989 | -105.2152
11.8026 | -105.2082
11.8063 | -105.204
11.8101 | -105.1981
11.8138 | -105.197
11.8175 | -105.1969
11.8213 | -105.1995
11.825 | -105.2053
11.8287 | -105.2123
11.8325 | -105.224
11.8362 | -105.2399
11.84 | -105.258
11.8437 | -105.2848
11.8474 | -105.3115
11.8512 | -105.348
11.8549 | -105.3734
11.8586 | -105.3997
11.8624 | -105.43
11.8661 | -105.4603
11.8698 | -105.4909
11.8736 | -105.5232
11.8773 | -105.5501
11.881 | -105.5725
11.8848 | -105.5896
11.8885 | -105.6082
11.8922 | -105.6279
11.896 | -105.6409
11.8997 | -105.6479
11.9034 | -105.6577
11.9072 | -105.6677
11.9109 | -105.671
11.9147 | -105.6732
11.9184 | -105.6737
11.9221 | -105.6752
11.9259 | -105.6711
11.9296 | -105.6625
11.9333 | -105.652
11.9371 | -105.6421
11.9408 | -105.6295
11.9445 | -105.6203
11.9483 | -105.6191
11.952 | -105.6149
11.9557 | -105.6128
11.9595 | -105.6119
11.9632 | -105.6073
11.9669 | -105.6046
11.9707 | -105.6057
11.9744 | -105.6028
11.9781 | -105.5996
11.9819 | -105.597
11.9856 | -105.5955
11.9894 | -105.5944
11.9931 | -105.5868
11.9968 | -105.5799
12.0006 | -105.5805
12.0043 | -105.5786
12.008 | -105.5737
12.0118 | -105.567
12.0155 | -105.5557
12.0192 | -105.547
12.023 | -105.5398
12.0267 | -105.5328
12.0304 | -105.5248
12.0342 | -105.5203
12.0379 | -105.5157
12.0416 | -105.5093
12.0454 | -105.5018
12.0491 | -105.497
12.0528 | -105.4932
12.0566 | -105.4909
12.0603 | -105.4891
12.0641 | -105.4842
12.0678 | -105.476
12.0715 | -105.4695
12.0753 | -105.4655
12.079 | -105.4598
12.0827 | -105.4554
12.0865 | -105.4493
12.0902 | -105.443
12.0939 | -105.4387
12.0977 | -105.4335
12.1014 | -105.4316
12.1051 | -105.4328
12.1089 | -105.434
12.1126 | -105.4358
12.1163 | -105.4362
12.1201 | -105.4313
12.1238 | -105.4273
12.1275 | -105.4246
12.1313 | -105.4205
12.135 | -105.4168
12.1388 | -105.4186
12.1425 | -105.421
12.1462 | -105.4256
12.15 | -105.4272
12.1537 | -105.4252
12.1574 | -105.4261
12.1612 | -105.4247
12.1649 | -105.4238
12.1686 | -105.4257
12.1724 | -105.4255
12.1761 | -105.4216
12.1798 | -105.4169
12.1836 | -105.4136
12.1873 | -105.41
12.191 | -105.4073
12.1948 | -105.4026
12.1985 | -105.3966
12.2022 | -105.3913
12.206 | -105.3876
12.2097 | -105.3797
12.2135 | -105.3698
12.2172 | -105.3647
12.2209 | -105.3624
12.2247 | -105.3604
12.2284 | -105.3575
12.2321 | -105.3529
12.2359 | -105.3516
12.2396 | -105.3525
12.2433 | -105.355
12.2471 | -105.3572
12.2508 | -105.3592
12.2545 | -105.3607
12.2583 | -105.3649
12.262 | -105.3717
12.2657 | -105.3769
12.2695 | -105.3852
12.2732 | -105.3931
12.2769 | -105.3994
12.2807 | -105.405
12.2844 | -105.4112
12.2882 | -105.4196
12.2919 | -105.4308
12.2956 | -105.4437
12.2994 | -105.4539
12.3031 | -105.4653
12.3068 | -105.4791
12.3106 | -105.4934
12.3143 | -105.5045
12.318 | -105.5145
12.3218 | -105.5252
12.3255 | -105.5413
12.3292 | -105.559
12.333 | -105.5704
12.3367 | -105.5833
12.3404 | -105.6006
12.3442 | -105.6158
12.3479 | -105.6261
12.3516 | -105.6347
12.3554 | -105.6464
12.3591 | -105.6584
12.3629 | -105.6711
12.3666 | -105.6833
12.3703 | -105.6915
12.3741 | -105.6962
12.3778 | -105.7007
12.3815 | -105.7066
12.3853 | -105.711
12.389 | -105.7145
12.3927 | -105.7191
12.3965 | -105.724
12.4002 | -105.7258
12.4039 | -105.7262
12.4077 | -105.7309
12.4114 | -105.7353
12.4151 | -105.7373
12.4189 | -105.7373
12.4226 | -105.7405
12.4263 | -105.7436
12.4301 | -105.7486
12.4338 | -105.7532
12.4376 | -105.7565
12.4413 | -105.7618
12.445 | -105.7674
12.4488 | -105.7713
12.4525 | -105.7741
12.4562 | -105.778
12.46 | -105.7826
12.4637 | -105.788
12.4674 | -105.7916
12.4712 | -105.7941
12.4749 | -105.7978
12.4786 | -105.8017
12.4824 | -105.7988
12.4861 | -105.7988
12.4898 | -105.7966
12.4936 | -105.7939
12.4973 | -105.7911
12.501 | -105.7874
12.5048 | -105.7824
12.5085 | -105.7769
12.5123 | -105.7716
12.516 | -105.7645
12.5197 | -105.753
12.5235 | -105.743
12.5272 | -105.737
12.5309 | -105.7297
12.5347 | -105.7188
12.5384 | -105.7075
12.5421 | -105.6969
12.5459 | -105.6822
12.5496 | -105.6675
12.5533 | -105.6548
12.5571 | -105.6438
12.5608 | -105.6298
12.5645 | -105.6178
12.5683 | -105.6062
12.572 | -105.5949
12.5758 | -105.5833
12.5795 | -105.5753
12.5832 | -105.5715
12.587 | -105.5688
12.5907 | -105.5663
12.5944 | -105.5664
12.5982 | -105.5671
12.6019 | -105.5701
12.6056 | -105.575
12.6094 | -105.5842
12.6131 | -105.5923
12.6168 | -105.6027
12.6206 | -105.613
12.6243 | -105.6213
12.628 | -105.6346
12.6318 | -105.6479
12.6355 | -105.6608
12.6392 | -105.6728
12.643 | -105.6831
12.6467 | -105.6925
12.6505 | -105.7023
12.6542 | -105.7102
12.6579 | -105.7163
12.6617 | -105.723
12.6654 | -105.7339
12.6691 | -105.7399
12.6729 | -105.7442
12.6766 | -105.7467
12.6803 | -105.7498
12.6841 | -105.75
12.6878 | -105.753
12.6915 | -105.7575
12.6953 | -105.7633
12.699 | -105.7687
12.7027 | -105.7725
12.7065 | -105.778
12.7102 | -105.7807
12.7139 | -105.7829
12.7177 | -105.784
12.7214 | -105.783
12.7252 | -105.7822
12.7289 | -105.7861
12.7326 | -105.7899
12.7364 | -105.7944
12.7401 | -105.7929
12.7438 | -105.7941
12.7476 | -105.7959
12.7513 | -105.7978
12.755 | -105.7994
12.7588 | -105.8021
12.7625 | -105.8013
12.7662 | -105.8006
12.77 | -105.8035
12.7737 | -105.8052
12.7774 | -105.81
12.7812 | -105.8157
12.7849 | -105.8188
12.7886 | -105.8234
12.7924 | -105.8299
12.7961 | -105.8391
12.7999 | -105.8506
12.8036 | -105.8621
12.8073 | -105.878
12.8111 | -105.8911
12.8148 | -105.904
12.8185 | -105.9218
12.8223 | -105.937
12.826 | -105.9473
12.8297 | -105.9596
12.8335 | -105.9729
12.8372 | -105.9887
12.8409 | -106.0052
12.8447 | -106.0163
12.8484 | -106.0253
12.8521 | -106.0357
12.8559 | -106.0458
12.8596 | -106.0528
12.8633 | -106.0568
12.8671 | -106.0594
12.8708 | -106.0582
12.8746 | -106.0553
12.8783 | -106.0522
12.882 | -106.0492
12.8858 | -106.0461
12.8895 | -106.0411
12.8932 | -106.0356
12.897 | -106.0315
12.9007 | -106.0267
12.9044 | -106.0218
12.9082 | -106.0187
12.9119 | -106.0132
12.9156 | -106.0079
12.9194 | -106.0041
12.9231 | -106.0077
12.9268 | -106.0089
12.9306 | -106.0154
12.9343 | -106.0146
12.938 | -106.0158
12.9418 | -106.0207
12.9455 | -106.0259
12.9493 | -106.0316
12.953 | -106.037
12.9567 | -106.0404
12.9605 | -106.0456
12.9642 | -106.048
12.9679 | -106.0482
12.9717 | -106.049
12.9754 | -106.0442
12.9791 | -106.0383
12.9829 | -106.0343
12.9866 | -106.0279
12.9903 | -106.0213
12.9941 | -106.012
12.9978 | -106.0025
13.0015 | -105.9927
13.0053 | -105.9799
13.009 | -105.9689
13.0127 | -105.9586
13.0165 | -105.9413
13.0202 | -105.923
13.024 | -105.9045
13.0277 | -105.8838
13.0314 | -105.866
13.0352 | -105.8484
13.0389 | -105.8299
13.0426 | -105.8152
13.0464 | -105.8045
13.0501 | -105.7927
13.0538 | -105.7826
13.0576 | -105.7748
13.0613 | -105.7698
13.065 | -105.7695
13.0688 | -105.7696
13.0725 | -105.7726
13.0762 | -105.7819
13.08 | -105.796
13.0837 | -105.8127
13.0874 | -105.8338
13.0912 | -105.8565
13.0949 | -105.8812
13.0987 | -105.9092
13.1024 | -105.9421
13.1061 | -105.9752
13.1099 | -106.0075
13.1136 | -106.0411
13.1173 | -106.0751
13.1211 | -106.1091
13.1248 | -106.1431
13.1285 | -106.1778
13.1323 | -106.2097
13.136 | -106.2391
13.1397 | -106.2702
13.1435 | -106.2987
13.1472 | -106.3214
13.1509 | -106.3407
13.1547 | -106.3563
13.1584 | -106.3696
13.1621 | -106.3792
13.1659 | -106.3876
13.1696 | -106.3894
13.1734 | -106.3842
13.1771 | -106.3755
13.1808 | -106.3657
13.1846 | -106.3511
13.1883 | -106.3346
13.192 | -106.3177
13.1958 | -106.2986
13.1995 | -106.2802
13.2032 | -106.2614
13.207 | -106.2424
13.2107 | -106.2229
13.2144 | -106.2048
13.2182 | -106.1879
13.2219 | -106.1721
13.2256 | -106.1582
13.2294 | -106.1471
13.2331 | -106.1354
13.2368 | -106.1265
13.2406 | -106.1249
13.2443 | -106.1254
13.2481 | -106.1269
13.2518 | -106.1323
13.2555 | -106.1383
13.2593 | -106.1499
13.263 | -106.1656
13.2667 | -106.1793
13.2705 | -106.198
13.2742 | -106.2189
13.2779 | -106.2396
13.2817 | -106.2637
13.2854 | -106.2864
13.2891 | -106.3124
13.2929 | -106.3361
13.2966 | -106.3595
13.3003 | -106.3829
13.3041 | -106.4071
13.3078 | -106.425
13.3115 | -106.4382
13.3153 | -106.4538
13.319 | -106.4657
13.3228 | -106.4735
13.3265 | -106.4787
13.3302 | -106.4779
13.334 | -106.4762
13.3377 | -106.4738
13.3414 | -106.4661
13.3452 | -106.4592
13.3489 | -106.4509
13.3526 | -106.4337
13.3564 | -106.414
13.3601 | -106.3953
13.3638 | -106.3772
13.3676 | -106.361
13.3713 | -106.3452
13.375 | -106.3291
13.3788 | -106.3118
13.3825 | -106.2932
13.3862 | -106.2758
13.39 | -106.2589
13.3937 | -106.2456
13.3975 | -106.2354
13.4012 | -106.2249
13.4049 | -106.213
13.4087 | -106.2022
13.4124 | -106.1935
13.4161 | -106.1894
13.4199 | -106.1824
13.4236 | -106.176
13.4273 | -106.1688
13.4311 | -106.1607
13.4348 | -106.1539
13.4385 | -106.1488
13.4423 | -106.1447
13.446 | -106.142
13.4497 | -106.139
13.4535 | -106.1342
13.4572 | -106.1322
13.4609 | -106.1275
13.4647 | -106.1225
13.4684 | -106.121
13.4722 | -106.118
13.4759 | -106.1146
13.4796 | -106.1122
13.4834 | -106.1052
13.4871 | -106.0988
13.4908 | -106.0935
13.4946 | -106.0897
13.4983 | -106.0864
13.502 | -106.0872
13.5058 | -106.0861
13.5095 | -106.0769
13.5132 | -106.0668
13.517 | -106.0599
13.5207 | -106.0554
13.5244 | -106.0504
13.5282 | -106.0467
13.5319 | -106.0441
13.5356 | -106.0431
13.5394 | -106.0436
13.5431 | -106.044
13.5469 | -106.0444
13.5506 | -106.0451
13.5543 | -106.0474
13.5581 | -106.0532
13.5618 | -106.0579
13.5655 | -106.0615
13.5693 | -106.0679
13.573 | -106.0781
13.5767 | -106.088
13.5805 | -106.1005
13.5842 | -106.1121
13.5879 | -106.1234
13.5917 | -106.1358
13.5954 | -106.1474
13.5991 | -106.1601
13.6029 | -106.176
13.6066 | -106.1863
13.6103 | -106.1944
13.6141 | -106.205
13.6178 | -106.2169
13.6216 | -106.2265
13.6253 | -106.2346
13.629 | -106.2417
13.6328 | -106.2459
13.6365 | -106.2474
13.6402 | -106.2462
13.644 | -106.2429
13.6477 | -106.2395
13.6514 | -106.235
13.6552 | -106.2306
13.6589 | -106.2254
13.6626 | -106.217
13.6664 | -106.2104
13.6701 | -106.2006
13.6738 | -106.1895
13.6776 | -106.1781
13.6813 | -106.1705
13.685 | -106.1613
13.6888 | -106.1494
13.6925 | -106.1391
13.6963 | -106.1282
13.7 | -106.1167
13.7037 | -106.1052
13.7075 | -106.0958
13.7112 | -106.0867
13.7149 | -106.0768
13.7187 | -106.0694
13.7224 | -106.0645
13.7261 | -106.0555
13.7299 | -106.0496
13.7336 | -106.0472
13.7373 | -106.0437
13.7411 | -106.0404
13.7448 | -106.0409
13.7485 | -106.0364
13.7523 | -106.0283
13.756 | -106.0195
13.7597 | -106.0139
13.7635 | -106.0093
13.7672 | -106.0035
13.771 | -105.9958
13.7747 | -105.9911
13.7784 | -105.9832
13.7822 | -105.9739
13.7859 | -105.9659
13.7896 | -105.9594
13.7934 | -105.9517
13.7971 | -105.9427
13.8008 | -105.9355
13.8046 | -105.9317
13.8083 | -105.9286
13.812 | -105.9249
13.8158 | -105.92
13.8195 | -105.9192
13.8232 | -105.9194
13.827 | -105.9204
13.8307 | -105.9226
13.8344 | -105.9248
13.8382 | -105.928
13.8419 | -105.9318
13.8457 | -105.9362
13.8494 | -105.9417
13.8531 | -105.9462
13.8569 | -105.9489
13.8606 | -105.9503
13.8643 | -105.9534
13.8681 | -105.9586
13.8718 | -105.9641
13.8755 | -105.9681
13.8793 | -105.9682
13.883 | -105.9702
13.8867 | -105.9736
13.8905 | -105.9729
13.8942 | -105.9694
13.8979 | -105.9685
13.9017 | -105.973
13.9054 | -105.9766
13.9091 | -105.978
13.9129 | -105.9771
13.9166 | -105.9744
13.9204 | -105.9734
13.9241 | -105.9746
13.9278 | -105.9739
13.9316 | -105.9713
13.9353 | -105.9699
13.939 | -105.9708
13.9428 | -105.9726
13.9465 | -105.9738
13.9502 | -105.9745
13.954 | -105.9735
13.9577 | -105.9704
13.9614 | -105.961
13.9652 | -105.9498
13.9689 | -105.9408
13.9726 | -105.9336
13.9764 | -105.9266
13.9801 | -105.9217
13.9838 | -105.9138
13.9876 | -105.9032
13.9913 | -105.8928
13.9951 | -105.8835
13.9988 | -105.8717
14.0025 | -105.8598
14.0063 | -105.8461
14.01 | -105.8324
14.0137 | -105.8228
14.0175 | -105.8146
14.0212 | -105.8042
14.0249 | -105.7949
14.0287 | -105.7856
14.0324 | -105.7799
14.0361 | -105.7781
14.0399 | -105.7781
14.0436 | -105.7768
14.0473 | -105.7768
14.0511 | -105.7783
14.0548 | -105.7828
14.0585 | -105.7879
14.0623 | -105.7954
14.066 | -105.8044
14.0698 | -105.8165
14.0735 | -105.8278
14.0772 | -105.8407
14.081 | -105.8566
14.0847 | -105.8715
14.0884 | -105.8889
14.0922 | -105.9061
14.0959 | -105.9246
14.0996 | -105.9409
14.1034 | -105.9536
14.1071 | -105.9656
14.1108 | -105.9794
14.1146 | -105.9956
14.1183 | -106.0117
14.122 | -106.0224
14.1258 | -106.0349
14.1295 | -106.0449
14.1332 | -106.0565
14.137 | -106.0681
14.1407 | -106.0796
14.1445 | -106.0931
14.1482 | -106.1062
14.1519 | -106.1196
14.1557 | -106.1324
14.1594 | -106.1461
14.1631 | -106.1632
14.1669 | -106.1789
14.1706 | -106.1902
14.1743 | -106.1986
14.1781 | -106.2098
14.1818 | -106.2192
14.1855 | -106.2285
14.1893 | -106.2371
14.193 | -106.2503
14.1967 | -106.2606
14.2005 | -106.2666
14.2042 | -106.274
14.2079 | -106.2826
14.2117 | -106.2881
14.2154 | -106.2949
14.2192 | -106.2983
14.2229 | -106.3045
14.2266 | -106.3114
14.2304 | -106.3177
14.2341 | -106.323
14.2378 | -106.3252
14.2416 | -106.3268
14.2453 | -106.3301
14.249 | -106.3367
14.2528 | -106.342
14.2565 | -106.3441
14.2602 | -106.3461
14.264 | -106.352
14.2677 | -106.3564
14.2714 | -106.3599
14.2752 | -106.3621
14.2789 | -106.3674
14.2826 | -106.373
14.2864 | -106.3762
14.2901 | -106.3828
14.2939 | -106.387
14.2976 | -106.392
14.3013 | -106.3963
14.3051 | -106.4009
14.3088 | -106.4046
14.3125 | -106.4064
14.3163 | -106.408
14.32 | -106.4041
14.3237 | -106.3995
14.3275 | -106.4016
14.3312 | -106.4028
14.3349 | -106.403
14.3387 | -106.4006
14.3424 | -106.3949
14.3461 | -106.3891
14.3499 | -106.3827
14.3536 | -106.3782
14.3573 | -106.3728
14.3611 | -106.3673
14.3648 | -106.3631
14.3686 | -106.3602
14.3723 | -106.3559
14.376 | -106.3463
14.3798 | -106.3377
14.3835 | -106.3295
14.3872 | -106.3221
14.391 | -106.3156
14.3947 | -106.3099
14.3984 | -106.3045
14.4022 | -106.2999
14.4059 | -106.2969
14.4096 | -106.292
14.4134 | -106.2882
14.4171 | -106.2848
14.4208 | -106.2788
14.4246 | -106.2709
14.4283 | -106.2595
14.432 | -106.2492
14.4358 | -106.2407
14.4395 | -106.2358
14.4433 | -106.2316
14.447 | -106.2254
14.4507 | -106.2179
14.4545 | -106.2094
14.4582 | -106.2031
14.4619 | -106.1973
14.4657 | -106.1911
14.4694 | -106.183
14.4731 | -106.175
14.4769 | -106.1648
14.4806 | -106.1569
14.4843 | -106.149
14.4881 | -106.1418
14.4918 | -106.1368
14.4955 | -106.1309
14.4993 | -106.1251
14.503 | -106.1221
14.5067 | -106.1212
14.5105 | -106.117
14.5142 | -106.1155
14.518 | -106.1175
14.5217 | -106.1166
14.5254 | -106.1188
14.5292 | -106.1198
14.5329 | -106.1188
14.5366 | -106.1194
14.5404 | -106.1243
14.5441 | -106.1294
14.5478 | -106.1414
14.5516 | -106.1492
14.5553 | -106.1606
14.559 | -106.1722
14.5628 | -106.1845
14.5665 | -106.195
14.5702 | -106.2065
14.574 | -106.2198
14.5777 | -106.2351
14.5814 | -106.2534
14.5852 | -106.2768
14.5889 | -106.2987
14.5927 | -106.3188
14.5964 | -106.3401
14.6001 | -106.359
14.6039 | -106.3773
14.6076 | -106.3935
14.6113 | -106.4109
14.6151 | -106.431
14.6188 | -106.4497
14.6225 | -106.4675
14.6263 | -106.4844
14.63 | -106.5014
14.6337 | -106.5193
14.6375 | -106.5349
14.6412 | -106.5498
14.6449 | -106.5623
14.6487 | -106.5706
14.6524 | -106.5803
14.6561 | -106.5907
14.6599 | -106.5986
14.6636 | -106.607
14.6674 | -106.6119
14.6711 | -106.6194
14.6748 | -106.6251
14.6786 | -106.627
14.6823 | -106.6296
14.686 | -106.6329
14.6898 | -106.6355
14.6935 | -106.6379
14.6972 | -106.6383
14.701 | -106.6375
14.7047 | -106.637
14.7084 | -106.6366
14.7122 | -106.6351
14.7159 | -106.632
14.7196 | -106.6306
14.7234 | -106.6246
14.7271 | -106.6192
14.7308 | -106.6167
14.7346 | -106.6135
14.7383 | -106.6111
14.7421 | -106.6087
14.7458 | -106.6047
14.7495 | -106.603
14.7533 | -106.6026
14.757 | -106.6018
14.7607 | -106.5995
14.7645 | -106.5992
14.7682 | -106.5945
14.7719 | -106.5924
14.7757 | -106.5946
14.7794 | -106.6006
14.7831 | -106.6048
14.7869 | -106.6093
14.7906 | -106.6145
14.7943 | -106.6186
14.7981 | -106.6225
14.8018 | -106.6286
14.8055 | -106.6357
14.8093 | -106.6408
14.813 | -106.6467
14.8168 | -106.6526
14.8205 | -106.659
14.8242 | -106.6635
14.828 | -106.663
14.8317 | -106.6615
14.8354 | -106.6592
14.8392 | -106.657
14.8429 | -106.6504
14.8466 | -106.6406
14.8504 | -106.6319
14.8541 | -106.6227
14.8578 | -106.6084
14.8616 | -106.594
14.8653 | -106.5781
14.869 | -106.5618
14.8728 | -106.5462
14.8765 | -106.5293
14.8802 | -106.5123
14.884 | -106.4912
14.8877 | -106.4733
14.8915 | -106.4604
14.8952 | -106.4462
14.8989 | -106.4311
14.9027 | -106.414
14.9064 | -106.3973
14.9101 | -106.384
14.9139 | -106.3734
14.9176 | -106.3629
14.9213 | -106.3538
14.9251 | -106.3449
14.9288 | -106.3375
14.9325 | -106.3323
14.9363 | -106.3297
14.94 | -106.3265
14.9437 | -106.3204
14.9475 | -106.3162
14.9512 | -106.3139
14.9549 | -106.3132
14.9587 | -106.312
14.9624 | -106.3093
14.9662 | -106.3108
14.9699 | -106.3154
14.9736 | -106.3209
14.9774 | -106.3257
14.9811 | -106.3335
14.9848 | -106.3411
14.9886 | -106.3455
14.9923 | -106.351
14.996 | -106.3607
14.9998 | -106.3707
15.0035 | -106.3797
15.0072 | -106.3909
15.011 | -106.4012
15.0147 | -106.4048
15.0184 | -106.4116
15.0222 | -106.4225
15.0259 | -106.4334
15.0296 | -106.4419
15.0334 | -106.4473
15.0371 | -106.4534
15.0409 | -106.4616
15.0446 | -106.4688
15.0483 | -106.4743
15.0521 | -106.4751
15.0558 | -106.4767
15.0595 | -106.4809
15.0633 | -106.4818
15.067 | -106.4802
15.0707 | -106.4793
15.0745 | -106.4775
15.0782 | -106.4752
15.0819 | -106.4726
15.0857 | -106.4633
15.0894 | -106.4569
15.0931 | -106.452
15.0969 | -106.4481
15.1006 | -106.4451
15.1043 | -106.4396
15.1081 | -106.4315
15.1118 | -106.4198
15.1156 | -106.4084
15.1193 | -106.4016
15.123 | -106.3953
15.1268 | -106.3879
15.1305 | -106.3794
15.1342 | -106.3732
15.138 | -106.3672
15.1417 | -106.3616
15.1454 | -106.3572
15.1492 | -106.355
15.1529 | -106.3532
15.1566 | -106.3529
15.1604 | -106.356
15.1641 | -106.3604
15.1678 | -106.3627
15.1716 | -106.366
15.1753 | -106.3739
15.179 | -106.3842
15.1828 | -106.3929
15.1865 | -106.4025
15.1903 | -106.4142
15.194 | -106.4306
15.1977 | -106.4445
15.2015 | -106.4568
15.2052 | -106.4729
15.2089 | -106.4895
15.2127 | -106.505
15.2164 | -106.5172
15.2201 | -106.5293
15.2239 | -106.5423
15.2276 | -106.5586
15.2313 | -106.5784
15.2351 | -106.5945
15.2388 | -106.6076
15.2425 | -106.6203
15.2463 | -106.6347
15.25 | -106.6485
15.2537 | -106.6598
15.2575 | -106.6706
15.2612 | -106.6796
15.265 | -106.6906
15.2687 | -106.7015
15.2724 | -106.7096
15.2762 | -106.7111
15.2799 | -106.7145
15.2836 | -106.718
15.2874 | -106.7184
15.2911 | -106.7186
15.2948 | -106.7151
15.2986 | -106.7099
15.3023 | -106.706
15.306 | -106.7018
15.3098 | -106.6964
15.3135 | -106.6869
15.3172 | -106.6767
15.321 | -106.6683
15.3247 | -106.6596
15.3284 | -106.6519
15.3322 | -106.6456
15.3359 | -106.6386
15.3397 | -106.6277
15.3434 | -106.6154
15.3471 | -106.6045
15.3509 | -106.5956
15.3546 | -106.5843
15.3583 | -106.5759
15.3621 | -106.5675
15.3658 | -106.5565
15.3695 | -106.5464
15.3733 | -106.5355
15.377 | -106.5254
15.3807 | -106.5173
15.3845 | -106.5074
15.3882 | -106.5021
15.3919 | -106.4956
15.3957 | -106.4891
15.3994 | -106.4825
15.4031 | -106.4769
15.4069 | -106.4724
15.4106 | -106.4682
15.4144 | -106.4608
15.4181 | -106.4586
15.4218 | -106.4569
15.4256 | -106.4518
15.4293 | -106.448
15.433 | -106.4437
15.4368 | -106.4425
15.4405 | -106.4411
15.4442 | -106.4386
15.448 | -106.4337
15.4517 | -106.4281
15.4554 | -106.4234
15.4592 | -106.42
15.4629 | -106.4167
15.4666 | -106.4135
15.4704 | -106.4114
15.4741 | -106.4109
15.4778 | -106.4113
15.4816 | -106.4106
15.4853 | -106.4094
15.4891 | -106.408
15.4928 | -106.4106
15.4965 | -106.4151
15.5003 | -106.4147
15.504 | -106.415
15.5077 | -106.4171
15.5115 | -106.4167
15.5152 | -106.4173
15.5189 | -106.4173
15.5227 | -106.412
15.5264 | -106.4073
15.5301 | -106.4029
15.5339 | -106.3993
15.5376 | -106.3959
15.5413 | -106.3899
15.5451 | -106.3834
15.5488 | -106.3769
15.5525 | -106.3719
15.5563 | -106.366
15.56 | -106.3609
15.5638 | -106.3566
15.5675 | -106.3507
15.5712 | -106.3447
15.575 | -106.3378
15.5787 | -106.3326
15.5824 | -106.3292
15.5862 | -106.3257
15.5899 | -106.3198
15.5936 | -106.3185
15.5974 | -106.3185
15.6011 | -106.3162
15.6048 | -106.315
15.6086 | -106.312
15.6123 | -106.3093
15.616 | -106.3067
15.6198 | -106.3058
15.6235 | -106.3032
15.6272 | -106.3028
15.631 | -106.3005
15.6347 | -106.2977
15.6385 | -106.2973
15.6422 | -106.2955
15.6459 | -106.2943
15.6497 | -106.2936
15.6534 | -106.292
15.6571 | -106.2911
15.6609 | -106.2883
15.6646 | -106.2862
15.6683 | -106.2823
15.6721 | -106.276
15.6758 | -106.2705
15.6795 | -106.2652
15.6833 | -106.2596
15.687 | -106.255
15.6907 | -106.2483
15.6945 | -106.2432
15.6982 | -106.2391
15.7019 | -106.2316
15.7057 | -106.2258
15.7094 | -106.2216
15.7132 | -106.2177
15.7169 | -106.2102
15.7206 | -106.2044
15.7244 | -106.2039
15.7281 | -106.2
15.7318 | -106.1967
15.7356 | -106.1952
15.7393 | -106.1912
15.743 | -106.1869
15.7468 | -106.1869
15.7505 | -106.1847
15.7542 | -106.1817
15.758 | -106.1765
15.7617 | -106.1698
15.7654 | -106.1626
15.7692 | -106.1563
15.7729 | -106.1499
15.7766 | -106.1395
15.7804 | -106.1283
15.7841 | -106.117
15.7879 | -106.1056
15.7916 | -106.0945
15.7953 | -106.08
15.7991 | -106.0666
15.8028 | -106.0502
15.8065 | -106.0346
15.8103 | -106.0202
15.814 | -106.0057
15.8177 | -105.9904
15.8215 | -105.9729
15.8252 | -105.9564
15.8289 | -105.942
15.8327 | -105.9276
15.8364 | -105.9155
15.8401 | -105.9039
15.8439 | -105.8924
15.8476 | -105.884
15.8513 | -105.8779
15.8551 | -105.8704
15.8588 | -105.8652
15.8626 | -105.8624
15.8663 | -105.8583
15.87 | -105.8572
15.8738 | -105.8594
15.8775 | -105.8637
15.8812 | -105.8671
15.885 | -105.8725
15.8887 | -105.8783
15.8924 | -105.8841
15.8962 | -105.8925
15.8999 | -105.9044
15.9036 | -105.9176
15.9074 | -105.9294
15.9111 | -105.9404
15.9148 | -105.9512
15.9186 | -105.963
15.9223 | -105.9754
15.926 | -105.9876
15.9298 | -106.004
15.9335 | -106.0225
15.9373 | -106.0387
15.941 | -106.0565
15.9447 | -106.0764
15.9485 | -106.0938
15.9522 | -106.1062
15.9559 | -106.1229
15.9597 | -106.1412
15.9634 | -106.1582
15.9671 | -106.1782
15.9709 | -106.1974
15.9746 | -106.2152
15.9783 | -106.2327
15.9821 | -106.2493
15.9858 | -106.2646
15.9895 | -106.2797
15.9933 | -106.2969
15.997 | -106.3132
16.0007 | -106.3261
16.0045 | -106.3417
16.0082 | -106.3571
16.012 | -106.3692
16.0157 | -106.3808
16.0194 | -106.3922
16.0232 | -106.405
16.0269 | -106.4165
16.0306 | -106.4284
16.0344 | -106.4368
16.0381 | -106.4457
16.0418 | -106.4544
16.0456 | -106.4615
16.0493 | -106.4682
16.053 | -106.475
16.0568 | -106.4823
16.0605 | -106.4874
16.0642 | -106.4923
16.068 | -106.4964
16.0717 | -106.5001
16.0754 | -106.5027
16.0792 | -106.502
16.0829 | -106.5022
16.0867 | -106.5045
16.0904 | -106.5062
16.0941 | -106.5063
16.0979 | -106.5022
16.1016 | -106.4996
16.1053 | -106.4961
16.1091 | -106.4927
16.1128 | -106.4863
16.1165 | -106.4822
16.1203 | -106.4765
16.124 | -106.4686
16.1277 | -106.4588
16.1315 | -106.4491
16.1352 | -106.4381
16.1389 | -106.4284
16.1427 | -106.4204
16.1464 | -106.4137
16.1501 | -106.4051
16.1539 | -106.3966
16.1576 | -106.3868
16.1614 | -106.3774
16.1651 | -106.3698
16.1688 | -106.3606
16.1726 | -106.3539
16.1763 | -106.3461
16.18 | -106.3394
16.1838 | -106.3383
16.1875 | -106.3359
16.1912 | -106.3326
16.195 | -106.3325
16.1987 | -106.3335
16.2024 | -106.3341
16.2062 | -106.3345
16.2099 | -106.3385
16.2136 | -106.345
16.2174 | -106.3541
16.2211 | -106.3631
16.2248 | -106.3707
16.2286 | -106.3795
16.2323 | -106.3955
16.2361 | -106.4118
16.2398 | -106.4277
16.2435 | -106.4446
16.2473 | -106.4607
16.251 | -106.4769
16.2547 | -106.494
16.2585 | -106.5092
16.2622 | -106.5267
16.2659 | -106.5481
16.2697 | -106.5665
16.2734 | -106.5847
16.2771 | -106.6048
16.2809 | -106.6239
16.2846 | -106.6406
16.2883 | -106.6558
16.2921 | -106.6724
16.2958 | -106.684
16.2995 | -106.6947
16.3033 | -106.7055
16.307 | -106.7175
16.3108 | -106.7303
16.3145 | -106.7376
16.3182 | -106.7416
16.322 | -106.7458
16.3257 | -106.7506
16.3294 | -106.754
16.3332 | -106.7524
16.3369 | -106.7506
16.3406 | -106.7513
16.3444 | -106.7525
16.3481 | -106.7502
16.3518 | -106.7459
16.3556 | -106.7419
16.3593 | -106.7357
16.363 | -106.7288
16.3668 | -106.7252
16.3705 | -106.7185
16.3742 | -106.711
16.378 | -106.7035
16.3817 | -106.6968
16.3855 | -106.6897
16.3892 | -106.6786
16.3929 | -106.6687
16.3967 | -106.6613
16.4004 | -106.6548
16.4041 | -106.6482
16.4079 | -106.6401
16.4116 | -106.6332
16.4153 | -106.6258
16.4191 | -106.6157
16.4228 | -106.6078
16.4265 | -106.6022
16.4303 | -106.5945
16.434 | -106.5869
16.4377 | -106.5844
16.4415 | -106.5845
16.4452 | -106.5837
16.4489 | -106.5822
16.4527 | -106.5831
16.4564 | -106.5841
16.4602 | -106.5846
16.4639 | -106.5831
16.4676 | -106.5829
16.4714 | -106.588
16.4751 | -106.5958
16.4788 | -106.6028
16.4826 | -106.6107
16.4863 | -106.6207
16.49 | -106.6305
16.4938 | -106.6427
16.4975 | -106.6529
16.5012 | -106.6619
16.505 | -106.6746
16.5087 | -106.6873
16.5124 | -106.6976
16.5162 | -106.7114
16.5199 | -106.724
16.5236 | -106.7347
16.5274 | -106.7439
16.5311 | -106.7568
16.5349 | -106.7669
16.5386 | -106.7758
16.5423 | -106.7781
16.5461 | -106.7788
16.5498 | -106.7854
16.5535 | -106.7935
16.5573 | -106.7983
16.561 | -106.8028
16.5647 | -106.8076
16.5685 | -106.8113
16.5722 | -106.8104
16.5759 | -106.8097
16.5797 | -106.8094
16.5834 | -106.8072
16.5871 | -106.8089
16.5909 | -106.8094
16.5946 | -106.8092
16.5983 | -106.8073
16.6021 | -106.8043
16.6058 | -106.8017
16.6096 | -106.7994
16.6133 | -106.7972
16.617 | -106.7944
16.6208 | -106.7906
16.6245 | -106.7917
16.6282 | -106.7889
16.632 | -106.787
16.6357 | -106.7836
16.6394 | -106.7775
16.6432 | -106.7731
16.6469 | -106.7686
16.6506 | -106.768
16.6544 | -106.7687
16.6581 | -106.7687
16.6618 | -106.768
16.6656 | -106.7669
16.6693 | -106.766
16.673 | -106.7643
16.6768 | -106.7626
16.6805 | -106.7645
16.6843 | -106.7673
16.688 | -106.7687
16.6917 | -106.7706
16.6955 | -106.7755
16.6992 | -106.7802
16.7029 | -106.7761
16.7067 | -106.7763
16.7104 | -106.7732
16.7141 | -106.77
16.7179 | -106.7684
16.7216 | -106.769
16.7253 | -106.7682
16.7291 | -106.7656
16.7328 | -106.7617
16.7365 | -106.7581
16.7403 | -106.7558
16.744 | -106.7501
16.7477 | -106.746
16.7515 | -106.7443
16.7552 | -106.7424
16.759 | -106.7422
16.7627 | -106.7411
16.7664 | -106.742
16.7702 | -106.739
16.7739 | -106.7377
16.7776 | -106.7359
16.7814 | -106.7383
16.7851 | -106.7412
16.7888 | -106.7381
16.7926 | -106.7346
16.7963 | -106.7349
16.8 | -106.7349
16.8038 | -106.7337
16.8075 | -106.7351
16.8112 | -106.7398
16.815 | -106.7414
16.8187 | -106.7414
16.8224 | -106.7459
16.8262 | -106.747
16.8299 | -106.7495
16.8337 | -106.7532
16.8374 | -106.7568
16.8411 | -106.7595
16.8449 | -106.7653
16.8486 | -106.7681
16.8523 | -106.7681
16.8561 | -106.7722
16.8598 | -106.7803
16.8635 | -106.7864
16.8673 | -106.7917
16.871 | -106.796
16.8747 | -106.8008
16.8785 | -106.8091
16.8822 | -106.8159
16.8859 | -106.8203
16.8897 | -106.8199
16.8934 | -106.8183
16.8971 | -106.816
16.9009 | -106.8151
16.9046 | -106.8149
16.9084 | -106.814
16.9121 | -106.8114
16.9158 | -106.809
16.9196 | -106.8056
16.9233 | -106.8022
16.927 | -106.798
16.9308 | -106.7949
16.9345 | -106.7917
16.9382 | -106.7868
16.942 | -106.7804
16.9457 | -106.7778
16.9494 | -106.772
16.9532 | -106.7661
16.9569 | -106.7596
16.9606 | -106.7548
16.9644 | -106.7466
16.9681 | -106.74
16.9718 | -106.7363
16.9756 | -106.7332
16.9793 | -106.7277
16.9831 | -106.7211
16.9868 | -106.7156
16.9905 | -106.7078
16.9943 | -106.6986
16.998 | -106.6908
17.0017 | -106.6848
17.0055 | -106.6782
17.0092 | -106.6708
17.0129 | -106.6663
17.0167 | -106.6616
17.0204 | -106.6548
17.0241 | -106.6493
17.0279 | -106.6449
17.0316 | -106.6418
17.0353 | -106.6395
17.0391 | -106.6353
17.0428 | -106.6321
17.0465 | -106.6298
17.0503 | -106.6288
17.054 | -106.6276
17.0578 | -106.6258
17.0615 | -106.6266
17.0652 | -106.6269
17.069 | -106.6301
17.0727 | -106.6328
17.0764 | -106.6363
17.0802 | -106.6431
17.0839 | -106.6491
17.0876 | -106.6527
17.0914 | -106.6586
17.0951 | -106.6696
17.0988 | -106.6784
17.1026 | -106.6861
17.1063 | -106.6941
17.11 | -106.702
17.1138 | -106.7111
17.1175 | -106.7218
17.1212 | -106.7317
17.125 | -106.7419
17.1287 | -106.7552
17.1325 | -106.7708
17.1362 | -106.7857
17.1399 | -106.8012
17.1437 | -106.8155
17.1474 | -106.8282
17.1511 | -106.8425
17.1549 | -106.857
17.1586 | -106.8693
17.1623 | -106.8828
17.1661 | -106.8963
17.1698 | -106.9097
17.1735 | -106.9213
17.1773 | -106.9309
17.181 | -106.9443
17.1847 | -106.9604
17.1885 | -106.9715
17.1922 | -106.9804
17.1959 | -106.9902
17.1997 | -107.0005
17.2034 | -107.0079
17.2072 | -107.0156
17.2109 | -107.0229
17.2146 | -107.0303
17.2184 | -107.0374
17.2221 | -107.0436
17.2258 | -107.0498
17.2296 | -107.0569
17.2333 | -107.0625
17.237 | -107.067
17.2408 | -107.0685
17.2445 | -107.0731
17.2482 | -107.0739
17.252 | -107.0757
17.2557 | -107.0792
17.2594 | -107.0792
17.2632 | -107.0788
17.2669 | -107.0793
17.2706 | -107.0762
17.2744 | -107.0725
17.2781 | -107.0701
17.2819 | -107.0663
17.2856 | -107.0632
17.2893 | -107.0595
17.2931 | -107.0555
17.2968 | -107.0488
17.3005 | -107.0426
17.3043 | -107.0361
17.308 | -107.0281
17.3117 | -107.0185
17.3155 | -107.009
17.3192 | -107.0002
17.3229 | -106.9928
17.3267 | -106.9839
17.3304 | -106.9778
17.3341 | -106.9718
17.3379 | -106.9637
17.3416 | -106.9549
17.3453 | -106.9444
17.3491 | -106.9334
17.3528 | -106.9266
17.3566 | -106.9222
17.3603 | -106.9127
17.364 | -106.9008
17.3678 | -106.8915
17.3715 | -106.8841
17.3752 | -106.8763
17.379 | -106.8677
17.3827 | -106.8599
17.3864 | -106.8546
17.3902 | -106.8508
17.3939 | -106.8478
17.3976 | -106.8432
17.4014 | -106.8414
17.4051 | -106.8392
17.4088 | -106.834
17.4126 | -106.8303
17.4163 | -106.8286
17.42 | -106.8294
17.4238 | -106.8318
17.4275 | -106.832
17.4313 | -106.8291
17.435 | -106.8312
17.4387 | -106.8333
17.4425 | -106.8341
17.4462 | -106.8351
17.4499 | -106.8376
17.4537 | -106.8396
17.4574 | -106.8385
17.4611 | -106.8362
17.4649 | -106.8334
17.4686 | -106.8312
17.4723 | -106.8304
17.4761 | -106.8285
17.4798 | -106.825
17.4835 | -106.8203
17.4873 | -106.8148
17.491 | -106.8082
17.4947 | -106.8006
17.4985 | -106.7944
17.5022 | -106.7894
17.506 | -106.7839
17.5097 | -106.7801
17.5134 | -106.7751
17.5172 | -106.7719
17.5209 | -106.7705
17.5246 | -106.7673
17.5284 | -106.7643
17.5321 | -106.7631
17.5358 | -106.7643
17.5396 | -106.7691
17.5433 | -106.7741
17.547 | -106.7771
17.5508 | -106.7813
17.5545 | -106.788
17.5582 | -106.7947
17.562 | -106.8028
17.5657 | -106.8113
17.5694 | -106.8194
17.5732 | -106.8276
17.5769 | -106.8375
17.5807 | -106.8486
17.5844 | -106.8587
17.5881 | -106.8712
17.5919 | -106.8826
17.5956 | -106.8966
17.5993 | -106.9089
17.6031 | -106.9179
17.6068 | -106.9288
17.6105 | -106.9416
17.6143 | -106.9526
17.618 | -106.9628
17.6217 | -106.9733
17.6255 | -106.9844
17.6292 | -106.9963
17.6329 | -107.0076
17.6367 | -107.0179
17.6404 | -107.0273
17.6441 | -107.036
17.6479 | -107.0467
17.6516 | -107.0551
17.6554 | -107.0642
17.6591 | -107.0734
17.6628 | -107.0808
17.6666 | -107.0879
17.6703 | -107.0924
17.674 | -107.0952
17.6778 | -107.0992
17.6815 | -107.1025
17.6852 | -107.1019
17.689 | -107.1035
17.6927 | -107.105
17.6964 | -107.1067
17.7002 | -107.1075
17.7039 | -107.1065
17.7076 | -107.1045
17.7114 | -107.1028
17.7151 | -107.1002
17.7188 | -107.0964
17.7226 | -107.0908
17.7263 | -107.0866
17.7301 | -107.0821
17.7338 | -107.0764
17.7375 | -107.0708
17.7413 | -107.0641
17.745 | -107.059
17.7487 | -107.0539
17.7525 | -107.0483
17.7562 | -107.0415
17.7599 | -107.0322
17.7637 | -107.0273
17.7674 | -107.021
17.7711 | -107.013
17.7749 | -107.0086
17.7786 | -107.0091
17.7823 | -107.0064
17.7861 | -107.0011
17.7898 | -106.995
17.7935 | -106.9893
17.7973 | -106.9864
17.801 | -106.985
17.8048 | -106.9832
17.8085 | -106.9822
17.8122 | -106.9813
17.816 | -106.9791
17.8197 | -106.9768
17.8234 | -106.9747
17.8272 | -106.9732
17.8309 | -106.9704
17.8346 | -106.9668
17.8384 | -106.9663
17.8421 | -106.9633
17.8458 | -106.9576
17.8496 | -106.9529
17.8533 | -106.9505
17.857 | -106.9459
17.8608 | -106.9385
17.8645 | -106.9324
17.8682 | -106.9283
17.872 | -106.9221
17.8757 | -106.9142
17.8795 | -106.905
17.8832 | -106.8958
17.8869 | -106.8882
17.8907 | -106.8798
17.8944 | -106.8721
17.8981 | -106.865
17.9019 | -106.8556
17.9056 | -106.8459
17.9093 | -106.8393
17.9131 | -106.8332
17.9168 | -106.8286
17.9205 | -106.8231
17.9243 | -106.8177
17.928 | -106.8127
17.9317 | -106.8086
17.9355 | -106.8045
17.9392 | -106.8002
17.9429 | -106.7959
17.9467 | -106.7916
17.9504 | -106.788
17.9542 | -106.7843
17.9579 | -106.783
17.9616 | -106.7843
17.9654 | -106.7855
17.9691 | -106.7873
17.9728 | -106.7897
17.9766 | -106.7907
17.9803 | -106.7909
17.984 | -106.7932
17.9878 | -106.7966
17.9915 | -106.8025
17.9952 | -106.8088
17.999 | -106.8117
18.0027 | -106.8148
18.0064 | -106.8182
18.0102 | -106.8237
18.0139 | -106.8317
18.0176 | -106.8373
18.0214 | -106.8421
18.0251 | -106.8473
18.0289 | -106.8515
18.0326 | -106.8551
18.0363 | -106.8586
18.0401 | -106.8621
18.0438 | -106.8667
18.0475 | -106.8719
18.0513 | -106.8767
18.055 | -106.8775
18.0587 | -106.8803
18.0625 | -106.8837
18.0662 | -106.8839
18.0699 | -106.8851
18.0737 | -106.887
18.0774 | -106.8888
18.0811 | -106.8919
18.0849 | -106.8942
18.0886 | -106.8955
18.0923 | -106.8982
18.0961 | -106.9
18.0998 | -106.9039
18.1036 | -106.9078
18.1073 | -106.91
18.111 | -106.9142
18.1148 | -106.9201
18.1185 | -106.9263
18.1222 | -106.9325
18.126 | -106.9388
18.1297 | -106.9452
18.1334 | -106.9521
18.1372 | -106.9583
18.1409 | -106.9631
18.1446 | -106.9706
18.1484 | -106.9775
18.1521 | -106.9828
18.1558 | -106.9913
18.1596 | -106.9997
18.1633 | -107.008
18.167 | -107.0096
18.1708 | -107.0125
18.1745 | -107.0153
18.1783 | -107.0181
18.182 | -107.022
18.1857 | -107.0257
18.1895 | -107.0265
18.1932 | -107.0271
18.1969 | -107.0293
18.2007 | -107.0298
18.2044 | -107.0288
18.2081 | -107.0276
18.2119 | -107.0235
18.2156 | -107.0179
18.2193 | -107.0107
18.2231 | -107.004
18.2268 | -106.9983
18.2305 | -106.9915
18.2343 | -106.983
18.238 | -106.9746
18.2417 | -106.9673
18.2455 | -106.9592
18.2492 | -106.9475
18.253 | -106.9378
18.2567 | -106.9313
18.2604 | -106.9262
18.2642 | -106.9188
18.2679 | -106.9112
18.2716 | -106.9044
18.2754 | -106.9003
18.2791 | -106.8985
18.2828 | -106.8951
18.2866 | -106.8924
18.2903 | -106.8917
18.294 | -106.8914
18.2978 | -106.8938
18.3015 | -106.8958
18.3052 | -106.8978
18.309 | -106.9011
18.3127 | -106.9069
18.3164 | -106.9128
18.3202 | -106.9238
18.3239 | -106.9315
18.3277 | -106.9415
18.3314 | -106.9484
18.3351 | -106.9552
18.3389 | -106.9637
18.3426 | -106.973
18.3463 | -106.9822
18.3501 | -106.9918
18.3538 | -107
18.3575 | -107.0094
18.3613 | -107.0178
18.365 | -107.0289
18.3687 | -107.0385
18.3725 | -107.0455
18.3762 | -107.0542
18.3799 | -107.0638
18.3837 | -107.0723
18.3874 | -107.0793
18.3911 | -107.0862
18.3949 | -107.0923
18.3986 | -107.1007
18.4024 | -107.1101
18.4061 | -107.1154
18.4098 | -107.12
18.4136 | -107.1273
18.4173 | -107.1324
18.421 | -107.1366
18.4248 | -107.1386
18.4285 | -107.1421
18.4322 | -107.1437
18.436 | -107.1439
18.4397 | -107.1428
18.4434 | -107.1405
18.4472 | -107.1389
18.4509 | -107.1368
18.4546 | -107.1357
18.4584 | -107.1295
18.4621 | -107.1241
18.4658 | -107.1195
18.4696 | -107.1132
18.4733 | -107.1086
18.4771 | -107.1047
18.4808 | -107.0999
18.4845 | -107.0937
18.4883 | -107.0865
18.492 | -107.0799
18.4957 | -107.0739
18.4995 | -107.0693
18.5032 | -107.062
18.5069 | -107.0561
18.5107 | -107.0504
18.5144 | -107.0447
18.5181 | -107.0398
18.5219 | -107.0331
18.5256 | -107.0266
18.5293 | -107.0197
18.5331 | -107.014
18.5368 | -107.0045
18.5405 | -106.9962
18.5443 | -106.9925
18.548 | -106.9875
18.5518 | -106.9798
18.5555 | -106.9726
18.5592 | -106.9659
18.563 | -106.9622
18.5667 | -106.9565
18.5704 | -106.9514
18.5742 | -106.9496
18.5779 | -106.945
18.5816 | -106.9387
18.5854 | -106.933
18.5891 | -106.9281
18.5928 | -106.9224
18.5966 | -106.92
18.6003 | -106.9158
18.604 | -106.9112
18.6078 | -106.9088
18.6115 | -106.9071
18.6152 | -106.9069
18.619 | -106.906
18.6227 | -106.9029
18.6265 | -106.9002
18.6302 | -106.8984
18.6339 | -106.897
18.6377 | -106.8949
18.6414 | -106.8933
18.6451 | -106.8932
18.6489 | -106.8935
18.6526 | -106.895
18.6563 | -106.8958
18.6601 | -106.8975
18.6638 | -106.8993
18.6675 | -106.9005
18.6713 | -106.901
18.675 | -106.9028
18.6787 | -106.9031
18.6825 | -106.9015
18.6862 | -106.9022
18.6899 | -106.9044
18.6937 | -106.9035
18.6974 | -106.9008
18.7012 | -106.8957
18.7049 | -106.8912
18.7086 | -106.8868
18.7124 | -106.8793
18.7161 | -106.8714
18.7198 | -106.8613
18.7236 | -106.8525
18.7273 | -106.8418
18.731 | -106.8285
18.7348 | -106.816
18.7385 | -106.8032
18.7422 | -106.7887
18.746 | -106.7742
18.7497 | -106.7597
18.7534 | -106.746
18.7572 | -106.7317
18.7609 | -106.7173
18.7646 | -106.7006
18.7684 | -106.6861
18.7721 | -106.6731
18.7759 | -106.6603
18.7796 | -106.6467
18.7833 | -106.6328
18.7871 | -106.6211
18.7908 | -106.6111
18.7945 | -106.6001
18.7983 | -106.5921
18.802 | -106.5883
18.8057 | -106.582
18.8095 | -106.5751
18.8132 | -106.5693
18.8169 | -106.5656
18.8207 | -106.5633
18.8244 | -106.5619
18.8281 | -106.5612
18.8319 | -106.5625
18.8356 | -106.5644
18.8393 | -106.5651
18.8431 | -106.5695
18.8468 | -106.5755
18.8506 | -106.5824
18.8543 | -106.5873
18.858 | -106.6006
18.8618 | -106.6108
18.8655 | -106.6213
18.8692 | -106.633
18.873 | -106.6438
18.8767 | -106.6534
18.8804 | -106.6632
18.8842 | -106.675
18.8879 | -106.6894
18.8916 | -106.7056
18.8954 | -106.7227
18.8991 | -106.7372
18.9028 | -106.7532
18.9066 | -106.7702
18.9103 | -106.7858
18.914 | -106.799
18.9178 | -106.8147
18.9215 | -106.8309
18.9253 | -106.8489
18.929 | -106.8667
18.9327 | -106.8838
18.9365 | -106.9001
18.9402 | -106.9164
18.9439 | -106.9349
18.9477 | -106.9527
18.9514 | -106.9704
18.9551 | -106.9874
18.9589 | -107.0043
18.9626 | -107.0238
18.9663 | -107.0426
18.9701 | -107.0608
18.9738 | -107.0764
18.9775 | -107.0916
18.9813 | -107.1068
18.985 | -107.1207
18.9887 | -107.1344
18.9925 | -107.1475
18.9962 | -107.1589
19 | -107.1704
19.0037 | -107.1778
19.0074 | -107.1848
19.0112 | -107.1923
19.0149 | -107.1987
19.0186 | -107.202
19.0224 | -107.2042
19.0261 | -107.2073
19.0298 | -107.2108
19.0336 | -107.2126
19.0373 | -107.2127
19.041 | -107.2101
19.0448 | -107.208
19.0485 | -107.2051
19.0522 | -107.2023
19.056 | -107.1979
19.0597 | -107.1949
19.0634 | -107.1943
19.0672 | -107.1882
19.0709 | -107.1791
19.0747 | -107.173
19.0784 | -107.1672
19.0821 | -107.1625
19.0859 | -107.1585
19.0896 | -107.1522
19.0933 | -107.1463
19.0971 | -107.1397
19.1008 | -107.1355
19.1045 | -107.1292
19.1083 | -107.1237
19.112 | -107.1197
19.1157 | -107.1157
19.1195 | -107.111
19.1232 | -107.1089
19.1269 | -107.1046
19.1307 | -107.1016
19.1344 | -107.0981
19.1381 | -107.0953
19.1419 | -107.0923
19.1456 | -107.091
19.1494 | -107.0901
19.1531 | -107.0915
19.1568 | -107.0915
19.1606 | -107.0907
19.1643 | -107.0887
19.168 | -107.0868
19.1718 | -107.0861
19.1755 | -107.0859
19.1792 | -107.0844
19.183 | -107.0839
19.1867 | -107.0847
19.1904 | -107.0839
19.1942 | -107.0833
19.1979 | -107.0819
19.2016 | -107.08
19.2054 | -107.0787
19.2091 | -107.0792
19.2128 | -107.0802
19.2166 | -107.0795
19.2203 | -107.0822
19.2241 | -107.0846
19.2278 | -107.0836
19.2315 | -107.0832
19.2353 | -107.0829
19.239 | -107.0843
19.2427 | -107.0888
19.2465 | -107.0943
19.2502 | -107.0978
19.2539 | -107.1008
19.2577 | -107.1048
19.2614 | -107.1099
19.2651 | -107.1163
19.2689 | -107.1208
19.2726 | -107.1245
19.2763 | -107.1303
19.2801 | -107.1348
19.2838 | -107.1389
19.2875 | -107.144
19.2913 | -107.1502
19.295 | -107.1574
19.2988 | -107.1632
19.3025 | -107.1678
19.3062 | -107.1725
19.31 | -107.1761
19.3137 | -107.1785
19.3174 | -107.1806
19.3212 | -107.1843
19.3249 | -107.189
19.3286 | -107.1928
19.3324 | -107.193
19.3361 | -107.1904
19.3398 | -107.189
19.3436 | -107.1899
19.3473 | -107.1911
19.351 | -107.1906
19.3548 | -107.1867
19.3585 | -107.1826
19.3622 | -107.1801
19.366 | -107.1778
19.3697 | -107.1765
19.3735 | -107.1734
19.3772 | -107.1719
19.3809 | -107.1691
19.3847 | -107.1651
19.3884 | -107.1603
19.3921 | -107.1563
19.3959 | -107.1501
19.3996 | -107.1444
19.4033 | -107.1406
19.4071 | -107.1352
19.4108 | -107.1281
19.4145 | -107.1183
19.4183 | -107.1127
19.422 | -107.1058
19.4257 | -107.0984
19.4295 | -107.0921
19.4332 | -107.0858
19.4369 | -107.0793
19.4407 | -107.0751
19.4444 | -107.0726
19.4482 | -107.0694
19.4519 | -107.0644
19.4556 | -107.0603
19.4594 | -107.0563
19.4631 | -107.0521
19.4668 | -107.049
19.4706 | -107.0466
19.4743 | -107.0446
19.478 | -107.0428
19.4818 | -107.0427
19.4855 | -107.0431
19.4892 | -107.0435
19.493 | -107.043
19.4967 | -107.0411
19.5004 | -107.0422
19.5042 | -107.0458
19.5079 | -107.049
19.5116 | -107.049
19.5154 | -107.0474
19.5191 | -107.0476
19.5229 | -107.0482
19.5266 | -107.0485
19.5303 | -107.0504
19.5341 | -107.0516
19.5378 | -107.0506
19.5415 | -107.0483
19.5453 | -107.0489
19.549 | -107.0509
19.5527 | -107.0526
19.5565 | -107.0527
19.5602 | -107.0493
19.5639 | -107.0471
19.5677 | -107.0458
19.5714 | -107.0435
19.5751 | -107.0407
19.5789 | -107.038
19.5826 | -107.0367
19.5863 | -107.0362
19.5901 | -107.0364
19.5938 | -107.0356
19.5976 | -107.0323
19.6013 | -107.0271
19.605 | -107.0225
19.6088 | -107.0196
19.6125 | -107.0177
19.6162 | -107.016
19.62 | -107.0117
19.6237 | -107.0045
19.6274 | -106.9985
19.6312 | -106.9936
19.6349 | -106.9888
19.6386 | -106.9848
19.6424 | -106.9818
19.6461 | -106.9769
19.6498 | -106.9717
19.6536 | -106.9668
19.6573 | -106.9633
19.661 | -106.9597
19.6648 | -106.9561
19.6685 | -106.9533
19.6723 | -106.9508
19.676 | -106.9483
19.6797 | -106.9453
19.6835 | -106.9391
19.6872 | -106.9345
19.6909 | -106.9324
19.6947 | -106.9327
19.6984 | -106.927
19.7021 | -106.9207
19.7059 | -106.9149
19.7096 | -106.909
19.7133 | -106.905
19.7171 | -106.9005
19.7208 | -106.8944
19.7245 | -106.8879
19.7283 | -106.8811
19.732 | -106.8753
19.7357 | -106.8707
19.7395 | -106.8665
19.7432 | -106.8607
19.747 | -106.8524
19.7507 | -106.842
19.7544 | -106.832
19.7582 | -106.8243
19.7619 | -106.8133
19.7656 | -106.8029
19.7694 | -106.7926
19.7731 | -106.7823
19.7768 | -106.7714
19.7806 | -106.7611
19.7843 | -106.7512
19.788 | -106.742
19.7918 | -106.7353
19.7955 | -106.7277
19.7992 | -106.7154
19.803 | -106.705
19.8067 | -106.6972
19.8104 | -106.6889
19.8142 | -106.6814
19.8179 | -106.6737
19.8217 | -106.6671
19.8254 | -106.6598
19.8291 | -106.6521
19.8329 | -106.6449
19.8366 | -106.6385
19.8403 | -106.636
19.8441 | -106.6316
19.8478 | -106.6272
19.8515 | -106.6221
19.8553 | -106.6191
19.859 | -106.6176
19.8627 | -106.6175
19.8665 | -106.6162
19.8702 | -106.6166
19.8739 | -106.6178
19.8777 | -106.6213
19.8814 | -106.6272
19.8851 | -106.6309
19.8889 | -106.6352
19.8926 | -106.6408
19.8964 | -106.6458
19.9001 | -106.654
19.9038 | -106.6631
19.9076 | -106.672
19.9113 | -106.6805
19.915 | -106.6889
19.9188 | -106.699
19.9225 | -106.7086
19.9262 | -106.7202
19.93 | -106.7326
19.9337 | -106.7459
19.9374 | -106.7566
19.9412 | -106.7683
19.9449 | -106.78
19.9486 | -106.7927
19.9524 | -106.8051
19.9561 | -106.8172
19.9598 | -106.8321
19.9636 | -106.8447
19.9673 | -106.8567
19.9711 | -106.8686
19.9748 | -106.879
19.9785 | -106.8878
19.9823 | -106.8985
19.986 | -106.9079
19.9897 | -106.9151
19.9935 | -106.9202
19.9972 | -106.9241
20.0009 | -106.9297
20.0047 | -106.9346
20.0084 | -106.936
20.0121 | -106.9378
20.0159 | -106.9393
20.0196 | -106.9396
20.0233 | -106.9396
20.0271 | -106.9413
20.0308 | -106.9425
20.0345 | -106.9446
20.0383 | -106.9445
20.042 | -106.9432
20.0458 | -106.9414
20.0495 | -106.9374
20.0532 | -106.9337
20.057 | -106.9302
20.0607 | -106.9274
20.0644 | -106.9232
20.0682 | -106.9196
20.0719 | -106.9136
20.0756 | -106.9085
20.0794 | -106.9033
20.0831 | -106.899
20.0868 | -106.8937
20.0906 | -106.8884
20.0943 | -106.8837
20.098 | -106.8805
20.1018 | -106.8767
20.1055 | -106.8704
20.1092 | -106.8643
20.113 | -106.8568
20.1167 | -106.85
20.1205 | -106.8447
20.1242 | -106.838
20.1279 | -106.8333
20.1317 | -106.8288
20.1354 | -106.8246
20.1391 | -106.8194
20.1429 | -106.8133
20.1466 | -106.8053
20.1503 | -106.8004
20.1541 | -106.7941
20.1578 | -106.7882
20.1615 | -106.7824
20.1653 | -106.7756
20.169 | -106.7699
20.1727 | -106.7621
20.1765 | -106.755
20.1802 | -106.749
20.1839 | -106.7403
20.1877 | -106.7319
20.1914 | -106.7235
20.1952 | -106.7159
20.1989 | -106.7094
20.2026 | -106.7006
20.2064 | -106.6924
20.2101 | -106.685
20.2138 | -106.6784
20.2176 | -106.6721
20.2213 | -106.6649
20.225 | -106.6587
20.2288 | -106.6534
20.2325 | -106.6482
20.2362 | -106.6431
20.24 | -106.6368
20.2437 | -106.6312
20.2474 | -106.6272
20.2512 | -106.6264
20.2549 | -106.6268
20.2586 | -106.625
20.2624 | -106.625
20.2661 | -106.6269
20.2699 | -106.6277
20.2736 | -106.6274
20.2773 | -106.6279
20.2811 | -106.6304
20.2848 | -106.6338
20.2885 | -106.6358
20.2923 | -106.6364
20.296 | -106.6363
20.2997 | -106.6387
20.3035 | -106.6429
20.3072 | -106.6467
20.3109 | -106.6486
20.3147 | -106.6493
20.3184 | -106.6509
20.3221 | -106.652
20.3259 | -106.654
20.3296 | -106.6549
20.3333 | -106.6553
20.3371 | -106.655
20.3408 | -106.6537
20.3446 | -106.653
20.3483 | -106.652
20.352 | -106.6497
20.3558 | -106.645
20.3595 | -106.6398
20.3632 | -106.634
20.367 | -106.63
20.3707 | -106.6255
20.3744 | -106.6192
20.3782 | -106.6112
20.3819 | -106.6037
20.3856 | -106.5945
20.3894 | -106.5867
20.3931 | -106.5804
20.3968 | -106.573
20.4006 | -106.5658
20.4043 | -106.5593
20.408 | -106.5504
20.4118 | -106.5392
20.4155 | -106.528
20.4193 | -106.5187
20.423 | -106.512
20.4267 | -106.5048
20.4305 | -106.498
20.4342 | -106.4925
20.4379 | -106.4866
20.4417 | -106.4833
20.4454 | -106.481
20.4491 | -106.4769
20.4529 | -106.4756
20.4566 | -106.4775
20.4603 | -106.4809
20.4641 | -106.4815
20.4678 | -106.4829
20.4715 | -106.4866
20.4753 | -106.4977
20.479 | -106.5038
20.4827 | -106.5124
20.4865 | -106.5203
20.4902 | -106.5273
20.494 | -106.5335
20.4977 | -106.5398
20.5014 | -106.547
20.5052 | -106.5547
20.5089 | -106.5615
20.5126 | -106.5696
20.5164 | -106.5765
20.5201 | -106.581
20.5238 | -106.5869
20.5276 | -106.5949
20.5313 | -106.6037
20.535 | -106.6109
20.5388 | -106.616
20.5425 | -106.6196
20.5462 | -106.6217
20.55 | -106.6238
20.5537 | -106.6271
20.5574 | -106.6296
20.5612 | -106.6326
20.5649 | -106.6367
20.5687 | -106.6377
20.5724 | -106.6348
20.5761 | -106.631
20.5799 | -106.6281
20.5836 | -106.6265
20.5873 | -106.6259
20.5911 | -106.6207
20.5948 | -106.6139
20.5985 | -106.6068
20.6023 | -106.6011
20.606 | -106.5948
20.6097 | -106.5886
20.6135 | -106.5826
20.6172 | -106.5761
20.6209 | -106.5681
20.6247 | -106.5597
20.6284 | -106.553
20.6321 | -106.5475
20.6359 | -106.5401
20.6396 | -106.5329
20.6434 | -106.5263
20.6471 | -106.5204
20.6508 | -106.5151
20.6546 | -106.5088
20.6583 | -106.5019
20.662 | -106.4968
20.6658 | -106.4913
20.6695 | -106.4836
20.6732 | -106.4783
20.677 | -106.4732
20.6807 | -106.4665
20.6844 | -106.461
20.6882 | -106.4563
20.6919 | -106.4537
20.6956 | -106.4488
20.6994 | -106.4452
20.7031 | -106.4426
20.7068 | -106.4374
20.7106 | -106.4306
20.7143 | -106.4236
20.7181 | -106.4161
20.7218 | -106.4091
20.7255 | -106.4016
20.7293 | -106.3952
20.733 | -106.3872
20.7367 | -106.3794
20.7405 | -106.3731
20.7442 | -106.3658
20.7479 | -106.357
20.7517 | -106.3474
20.7554 | -106.3393
20.7591 | -106.3302
20.7629 | -106.3213
20.7666 | -106.3099
20.7703 | -106.2984
20.7741 | -106.2878
20.7778 | -106.2772
20.7815 | -106.2686
20.7853 | -106.2591
20.789 | -106.2514
20.7928 | -106.2446
20.7965 | -106.2349
20.8002 | -106.2275
20.804 | -106.2214
20.8077 | -106.2143
20.8114 | -106.2078
20.8152 | -106.2009
20.8189 | -106.1942
20.8226 | -106.1862
20.8264 | -106.1795
20.8301 | -106.1743
20.8338 | -106.1681
20.8376 | -106.165
20.8413 | -106.1628
20.845 | -106.1602
20.8488 | -106.1542
20.8525 | -106.1473
20.8562 | -106.1409
20.86 | -106.1363
20.8637 | -106.1338
20.8675 | -106.1307
20.8712 | -106.1263
20.8749 | -106.1221
20.8787 | -106.1161
20.8824 | -106.1101
20.8861 | -106.1043
20.8899 | -106.0988
20.8936 | -106.0935
20.8973 | -106.0862
20.9011 | -106.079
20.9048 | -106.0739
20.9085 | -106.0667
20.9123 | -106.06
20.916 | -106.0553
20.9197 | -106.0512
20.9235 | -106.0449
20.9272 | -106.0381
20.9309 | -106.0324
20.9347 | -106.0259
20.9384 | -106.0197
20.9422 | -106.0148
20.9459 | -106.0113
20.9496 | -106.0064
20.9534 | -106.0002
20.9571 | -105.9937
20.9608 | -105.9884
20.9646 | -105.9848
20.9683 | -105.9799
20.972 | -105.9763
20.9758 | -105.9743
20.9795 | -105.9723
20.9832 | -105.9688
20.987 | -105.9662
20.9907 | -105.9652
20.9944 | -105.9662
20.9982 | -105.9688
21.0019 | -105.9716
21.0056 | -105.9735
21.0094 | -105.976
21.0131 | -105.985
21.0169 | -105.9902
21.0206 | -105.9952
21.0243 | -106.0002
21.0281 | -106.0059
21.0318 | -106.0137
21.0355 | -106.021
21.0393 | -106.03
21.043 | -106.0403
21.0467 | -106.0512
21.0505 | -106.0636
21.0542 | -106.0754
21.0579 | -106.0871
21.0617 | -106.0976
21.0654 | -106.1106
21.0691 | -106.1261
21.0729 | -106.1392
21.0766 | -106.151
21.0803 | -106.1629
21.0841 | -106.1764
21.0878 | -106.187
21.0916 | -106.1975
21.0953 | -106.2108
21.099 | -106.2219
21.1028 | -106.2323
21.1065 | -106.2436
21.1102 | -106.2548
21.114 | -106.2658
21.1177 | -106.2747
21.1214 | -106.2838
21.1252 | -106.2925
21.1289 | -106.2992
21.1326 | -106.3074
21.1364 | -106.3117
21.1401 | -106.3156
21.1438 | -106.3189
21.1476 | -106.322
21.1513 | -106.3278
21.155 | -106.3339
21.1588 | -106.3376
21.1625 | -106.3407
21.1663 | -106.3445
21.17 | -106.3471
21.1737 | -106.348
21.1775 | -106.3515
21.1812 | -106.3537
21.1849 | -106.3546
21.1887 | -106.3554
21.1924 | -106.3558
21.1961 | -106.3553
21.1999 | -106.3542
21.2036 | -106.3522
21.2073 | -106.352
21.2111 | -106.3513
21.2148 | -106.3499
21.2185 | -106.3498
21.2223 | -106.35
21.226 | -106.3498
21.2297 | -106.3505
21.2335 | -106.3507
21.2372 | -106.3518
21.241 | -106.3524
21.2447 | -106.3539
21.2484 | -106.3588
21.2522 | -106.3639
21.2559 | -106.3694
21.2596 | -106.3738
21.2634 | -106.3776
21.2671 | -106.382
21.2708 | -106.3862
21.2746 | -106.3905
21.2783 | -106.3944
21.282 | -106.4004
21.2858 | -106.4046
21.2895 | -106.4095
21.2932 | -106.4156
21.297 | -106.4204
21.3007 | -106.4277
21.3044 | -106.4365
21.3082 | -106.4442
21.3119 | -106.4504
21.3157 | -106.4559
21.3194 | -106.4617
21.3231 | -106.4703
21.3269 | -106.4792
21.3306 | -106.4855
21.3343 | -106.4917
21.3381 | -106.4987
21.3418 | -106.5045
21.3455 | -106.5091
21.3493 | -106.5165
21.353 | -106.5255
21.3567 | -106.5359
21.3605 | -106.5459
21.3642 | -106.5548
21.3679 | -106.5606
21.3717 | -106.5648
21.3754 | -106.5709
21.3791 | -106.5776
21.3829 | -106.5845
21.3866 | -106.5929
21.3904 | -106.6023
21.3941 | -106.6114
21.3978 | -106.617
21.4016 | -106.6232
21.4053 | -106.6305
21.409 | -106.6375
21.4128 | -106.6432
21.4165 | -106.6496
21.4202 | -106.6562
21.424 | -106.6625
21.4277 | -106.6687
21.4314 | -106.6757
21.4352 | -106.681
21.4389 | -106.6878
21.4426 | -106.6952
21.4464 | -106.7031
21.4501 | -106.712
21.4538 | -106.717
21.4576 | -106.7196
21.4613 | -106.7225
21.4651 | -106.7287
21.4688 | -106.7342
21.4725 | -106.7407
21.4763 | -106.7475
21.48 | -106.7526
21.4837 | -106.7599
21.4875 | -106.767
21.4912 | -106.7707
21.4949 | -106.7753
21.4987 | -106.7833
21.5024 | -106.7906
21.5061 | -106.7961
21.5099 | -106.8004
21.5136 | -106.8036
21.5173 | -106.8082
21.5211 | -106.8123
21.5248 | -106.8161
21.5285 | -106.8202
21.5323 | -106.8223
21.536 | -106.8264
21.5398 | -106.8264
21.5435 | -106.8244
21.5472 | -106.8225
21.551 | -106.8144
21.5547 | -106.8095
21.5584 | -106.8057
21.5622 | -106.8002
21.5659 | -106.7948
21.5696 | -106.7924
21.5734 | -106.7876
21.5771 | -106.7812
21.5808 | -106.7756
21.5846 | -106.7692
21.5883 | -106.7613
21.592 | -106.7529
21.5958 | -106.7432
21.5995 | -106.7318
21.6032 | -106.7214
21.607 | -106.7095
21.6107 | -106.6965
21.6145 | -106.6846
21.6182 | -106.6738
21.6219 | -106.6623
21.6257 | -106.6504
21.6294 | -106.635
21.6331 | -106.6192
21.6369 | -106.6056
21.6406 | -106.5913
21.6443 | -106.5783
21.6481 | -106.5674
21.6518 | -106.5568
21.6555 | -106.5447
21.6593 | -106.5304
21.663 | -106.5163
21.6667 | -106.5056
21.6705 | -106.4972
21.6742 | -106.49
21.6779 | -106.4831
21.6817 | -106.4763
21.6854 | -106.4712
21.6892 | -106.4651
21.6929 | -106.4596
21.6966 | -106.4554
21.7004 | -106.4517
21.7041 | -106.4514
21.7078 | -106.4521
21.7116 | -106.4525
21.7153 | -106.4532
21.719 | -106.4562
21.7228 | -106.4602
21.7265 | -106.4644
21.7302 | -106.4686
21.734 | -106.4724
21.7377 | -106.4782
21.7414 | -106.4874
21.7452 | -106.4961
21.7489 | -106.5061
21.7526 | -106.516
21.7564 | -106.5259
21.7601 | -106.5353
21.7639 | -106.545
21.7676 | -106.5545
21.7713 | -106.5619
21.7751 | -106.5719
21.7788 | -106.5808
21.7825 | -106.5881
21.7863 | -106.597
21.79 | -106.6046
21.7937 | -106.6113
21.7975 | -106.6205
21.8012 | -106.6269
21.8049 | -106.6316
21.8087 | -106.6348
21.8124 | -106.6384
21.8161 | -106.6431
21.8199 | -106.6467
21.8236 | -106.6466
21.8273 | -106.6454
21.8311 | -106.644
21.8348 | -106.6406
21.8386 | -106.6369
21.8423 | -106.6315
21.846 | -106.6286
21.8498 | -106.6268
21.8535 | -106.6218
21.8572 | -106.6166
21.861 | -106.6098
21.8647 | -106.6006
21.8684 | -106.591
21.8722 | -106.5815
21.8759 | -106.5718
21.8796 | -106.5618
21.8834 | -106.5501
21.8871 | -106.5378
21.8908 | -106.5239
21.8946 | -106.5107
21.8983 | -106.4998
21.902 | -106.4903
21.9058 | -106.4791
21.9095 | -106.469
21.9133 | -106.4585
21.917 | -106.4483
21.9207 | -106.4391
21.9245 | -106.43
21.9282 | -106.4208
21.9319 | -106.4121
21.9357 | -106.4042
21.9394 | -106.3966
21.9431 | -106.3893
21.9469 | -106.3816
21.9506 | -106.3757
21.9543 | -106.3679
21.9581 | -106.3589
21.9618 | -106.3493
21.9655 | -106.3392
21.9693 | -106.3329
21.973 | -106.3254
21.9767 | -106.3173
21.9805 | -106.3101
21.9842 | -106.3055
21.988 | -106.2994
21.9917 | -106.2903
21.9954 | -106.2815
21.9992 | -106.2723
22.0029 | -106.2648
22.0066 | -106.2603
22.0104 | -106.2565
22.0141 | -106.2515
22.0178 | -106.2455
22.0216 | -106.2392
22.0253 | -106.2338
22.029 | -106.2296
22.0328 | -106.2263
22.0365 | -106.2236
22.0402 | -106.2209
22.044 | -106.2199
22.0477 | -106.2174
22.0514 | -106.2132
22.0552 | -106.2108
22.0589 | -106.2088
22.0627 | -106.2066
22.0664 | -106.204
22.0701 | -106.2028
22.0739 | -106.2027
22.0776 | -106.2037
22.0813 | -106.2037
22.0851 | -106.2022
22.0888 | -106.2031
22.0925 | -106.2068
22.0963 | -106.2086
22.1 | -106.2077
22.1037 | -106.2068
22.1075 | -106.2087
22.1112 | -106.2086
22.1149 | -106.2096
22.1187 | -106.2115
22.1224 | -106.2137
22.1261 | -106.216
22.1299 | -106.2176
22.1336 | -106.2189
22.1374 | -106.2214
22.1411 | -106.2237
22.1448 | -106.2261
22.1486 | -106.2274
22.1523 | -106.2287
22.156 | -106.231
22.1598 | -106.2327
22.1635 | -106.235
22.1672 | -106.2377
22.171 | -106.2385
22.1747 | -106.2384
22.1784 | -106.2395
22.1822 | -106.2407
22.1859 | -106.2413
22.1896 | -106.2428
22.1934 | -106.2422
22.1971 | -106.2423
22.2008 | -106.2435
22.2046 | -106.2451
22.2083 | -106.2456
22.2121 | -106.2421
22.2158 | -106.2373
22.2195 | -106.2345
22.2233 | -106.2307
22.227 | -106.2291
22.2307 | -106.2266
22.2345 | -106.2233
22.2382 | -106.2183
22.2419 | -106.2146
22.2457 | -106.212
22.2494 | -106.2076
22.2531 | -106.2047
22.2569 | -106.2023
22.2606 | -106.1997
22.2643 | -106.1961
22.2681 | -106.193
22.2718 | -106.1901
22.2755 | -106.1876
22.2793 | -106.1856
22.283 | -106.1821
22.2868 | -106.1812
22.2905 | -106.1811
22.2942 | -106.1789
22.298 | -106.1756
22.3017 | -106.1746
22.3054 | -106.1763
22.3092 | -106.1766
22.3129 | -106.1791
22.3166 | -106.1798
22.3204 | -106.179
22.3241 | -106.1791
22.3278 | -106.1811
22.3316 | -106.1821
22.3353 | -106.1837
22.339 | -106.1861
22.3428 | -106.1882
22.3465 | -106.1897
22.3502 | -106.1931
22.354 | -106.1983
22.3577 | -106.2018
22.3615 | -106.2071
22.3652 | -106.213
22.3689 | -106.2193
22.3727 | -106.2256
22.3764 | -106.2318
22.3801 | -106.2397
22.3839 | -106.2472
22.3876 | -106.2544
22.3913 | -106.2611
22.3951 | -106.2696
22.3988 | -106.2775
22.4025 | -106.285
22.4063 | -106.2906
22.41 | -106.2954
22.4137 | -106.3023
22.4175 | -106.3085
22.4212 | -106.3127
22.4249 | -106.3162
22.4287 | -106.3213
22.4324 | -106.3259
22.4362 | -106.3294
22.4399 | -106.3343
22.4436 | -106.3383
22.4474 | -106.3409
22.4511 | -106.3447
22.4548 | -106.3465
22.4586 | -106.3492
22.4623 | -106.3508
22.466 | -106.3539
22.4698 | -106.3548
22.4735 | -106.3546
22.4772 | -106.3554
22.481 | -106.355
22.4847 | -106.3544
22.4884 | -106.3556
22.4922 | -106.3584
22.4959 | -106.361
22.4996 | -106.361
22.5034 | -106.3609
22.5071 | -106.3604
22.5109 | -106.3605
22.5146 | -106.3578
22.5183 | -106.355
22.5221 | -106.3549
22.5258 | -106.3549
22.5295 | -106.3554
22.5333 | -106.3552
22.537 | -106.354
22.5407 | -106.3525
22.5445 | -106.3496
22.5482 | -106.3482
22.5519 | -106.3487
22.5557 | -106.3502
22.5594 | -106.3509
22.5631 | -106.3488
22.5669 | -106.3467
22.5706 | -106.3461
22.5743 | -106.3415
22.5781 | -106.3343
22.5818 | -106.3281
22.5856 | -106.3225
22.5893 | -106.3177
22.593 | -106.3129
22.5968 | -106.3042
22.6005 | -106.2939
22.6042 | -106.2815
22.608 | -106.269
22.6117 | -106.259
22.6154 | -106.2492
22.6192 | -106.238
22.6229 | -106.225
22.6266 | -106.2123
22.6304 | -106.1906
22.6341 | -106.1746
22.6378 | -106.1583
22.6416 | -106.1434
22.6453 | -106.1271
22.649 | -106.1104
22.6528 | -106.0945
22.6565 | -106.0793
22.6603 | -106.0629
22.664 | -106.0445
22.6677 | -106.0255
22.6715 | -106.0046
22.6752 | -105.9854
22.6789 | -105.9683
22.6827 | -105.9493
22.6864 | -105.9313
22.6901 | -105.9139
22.6939 | -105.8951
22.6976 | -105.8794
22.7013 | -105.8648
22.7051 | -105.8524
22.7088 | -105.8411
22.7125 | -105.8307
22.7163 | -105.8204
22.72 | -105.8124
22.7237 | -105.8049
22.7275 | -105.8014
22.7312 | -105.8019
22.735 | -105.8018
22.7387 | -105.806
22.7424 | -105.8125
22.7462 | -105.8194
22.7499 | -105.8266
22.7536 | -105.8352
22.7574 | -105.846
22.7611 | -105.858
22.7648 | -105.8712
22.7686 | -105.8872
22.7723 | -105.9062
22.776 | -105.926
22.7798 | -105.9473
22.7835 | -105.9706
22.7872 | -105.9947
22.791 | -106.0195
22.7947 | -106.0446
22.7984 | -106.0705
22.8022 | -106.0963
22.8059 | -106.1232
22.8097 | -106.1508
22.8134 | -106.1799
22.8171 | -106.2094
22.8209 | -106.2405
22.8246 | -106.27
22.8283 | -106.3003
22.8321 | -106.3303
22.8358 | -106.3628
22.8395 | -106.3937
22.8433 | -106.4199
22.847 | -106.4468
22.8507 | -106.4743
22.8545 | -106.5002
22.8582 | -106.5254
22.8619 | -106.5495
22.8657 | -106.5713
22.8694 | -106.5915
22.8731 | -106.6094
22.8769 | -106.6271
22.8806 | -106.6427
22.8844 | -106.6609
22.8881 | -106.6786
22.8918 | -106.6918
22.8956 | -106.7069
22.8993 | -106.7228
22.903 | -106.7371
22.9068 | -106.7497
22.9105 | -106.7597
22.9142 | -106.7676
22.918 | -106.7786
22.9217 | -106.7866
22.9254 | -106.792
22.9292 | -106.7957
22.9329 | -106.7968
22.9366 | -106.7964
22.9404 | -106.7974
22.9441 | -106.7992
22.9478 | -106.8015
22.9516 | -106.8042
22.9553 | -106.8057
22.9591 | -106.8073
22.9628 | -106.808
22.9665 | -106.8067
22.9703 | -106.8069
22.974 | -106.8085
22.9777 | -106.8091
22.9815 | -106.8094
22.9852 | -106.8104
22.9889 | -106.813
22.9927 | -106.8165
22.9964 | -106.82
23.0001 | -106.821
23.0039 | -106.8234
23.0076 | -106.8287
23.0113 | -106.8342
23.0151 | -106.8393
23.0188 | -106.8472
23.0225 | -106.8532
23.0263 | -106.8564
23.03 | -106.8599
23.0338 | -106.863
23.0375 | -106.8662
23.0412 | -106.8692
23.045 | -106.8722
23.0487 | -106.8787
23.0524 | -106.8832
23.0562 | -106.8866
23.0599 | -106.8902
23.0636 | -106.892
23.0674 | -106.8957
23.0711 | -106.9003
23.0748 | -106.9051
23.0786 | -106.909
23.0823 | -106.9093
23.086 | -106.9121
23.0898 | -106.9143
23.0935 | -106.9156
23.0972 | -106.9156
23.101 | -106.9175
23.1047 | -106.9173
23.1085 | -106.9165
23.1122 | -106.916
23.1159 | -106.9169
23.1197 | -106.9158
23.1234 | -106.9125
23.1271 | -106.9093
23.1309 | -106.9053
23.1346 | -106.9003
23.1383 | -106.8934
23.1421 | -106.8865
23.1458 | -106.8798
23.1495 | -106.8745
23.1533 | -106.8676
23.157 | -106.8603
23.1607 | -106.8531
23.1645 | -106.8454
23.1682 | -106.8381
23.1719 | -106.8316
23.1757 | -106.8239
23.1794 | -106.8171
23.1832 | -106.8105
23.1869 | -106.8035
23.1906 | -106.7972
23.1944 | -106.7898
23.1981 | -106.7813
23.2018 | -106.7738
23.2056 | -106.7653
23.2093 | -106.7562
23.213 | -106.7471
23.2168 | -106.7393
23.2205 | -106.7327
23.2242 | -106.7263
23.228 | -106.7202
23.2317 | -106.7145
23.2354 | -106.7087
23.2392 | -106.703
23.2429 | -106.6985
23.2466 | -106.6948
23.2504 | -106.692
23.2541 | -106.6913
23.2579 | -106.6901
23.2616 | -106.6905
23.2653 | -106.6915
23.2691 | -106.6918
23.2728 | -106.694
23.2765 | -106.6954
23.2803 | -106.6978
23.284 | -106.7034
23.2877 | -106.7078
23.2915 | -106.7124
23.2952 | -106.7188
23.2989 | -106.7258
23.3027 | -106.7322
23.3064 | -106.7416
23.3101 | -106.7506
23.3139 | -106.7582
23.3176 | -106.7672
23.3213 | -106.7783
23.3251 | -106.7909
23.3288 | -106.803
23.3326 | -106.8148
23.3363 | -106.8263
23.34 | -106.8351
23.3438 | -106.8454
23.3475 | -106.8564
23.3512 | -106.8692
23.355 | -106.8827
23.3587 | -106.8953
23.3624 | -106.9098
23.3662 | -106.9235
23.3699 | -106.9372
23.3736 | -106.9504
23.3774 | -106.9613
23.3811 | -106.9722
23.3848 | -106.9855
23.3886 | -106.998
23.3923 | -107.0102
23.396 | -107.0206
23.3998 | -107.0287
23.4035 | -107.0377
23.4073 | -107.049
23.411 | -107.0611
23.4147 | -107.0698
23.4185 | -107.0799
23.4222 | -107.0906
23.4259 | -107.0987
23.4297 | -107.1042
23.4334 | -107.1108
23.4371 | -107.1178
23.4409 | -107.1239
23.4446 | -107.1299
23.4483 | -107.1362
23.4521 | -107.1404
23.4558 | -107.1455
23.4595 | -107.1497
23.4633 | -107.1536
23.467 | -107.1571
23.4707 | -107.1589
23.4745 | -107.1575
23.4782 | -107.1547
23.482 | -107.1532
23.4857 | -107.1513
23.4894 | -107.1494
23.4932 | -107.1465
23.4969 | -107.1405
23.5006 | -107.1361
23.5044 | -107.1337
23.5081 | -107.1309
23.5118 | -107.1266
23.5156 | -107.1222
23.5193 | -107.1169
23.523 | -107.1116
23.5268 | -107.1062
23.5305 | -107.1013
23.5342 | -107.095
23.538 | -107.0903
23.5417 | -107.0856
23.5454 | -107.0795
23.5492 | -107.0711
23.5529 | -107.0638
23.5567 | -107.0577
23.5604 | -107.0508
23.5641 | -107.044
23.5679 | -107.0387
23.5716 | -107.0343
23.5753 | -107.0271
23.5791 | -107.0199
23.5828 | -107.0152
23.5865 | -107.0109
23.5903 | -107.0058
23.594 | -106.9999
23.5977 | -106.9957
23.6015 | -106.9933
23.6052 | -106.9907
23.6089 | -106.9886
23.6127 | -106.9869
23.6164 | -106.9867
23.6201 | -106.9871
23.6239 | -106.9886
23.6276 | -106.9904
23.6314 | -106.9937
23.6351 | -106.997
23.6388 | -106.9998
23.6426 | -107.0047
23.6463 | -107.0122
23.65 | -107.019
23.6538 | -107.0247
23.6575 | -107.0309
23.6612 | -107.038
23.665 | -107.0443
23.6687 | -107.0527
23.6724 | -107.0645
23.6762 | -107.0749
23.6799 | -107.0838
23.6836 | -107.0927
23.6874 | -107.1037
23.6911 | -107.114
23.6948 | -107.1232
23.6986 | -107.1344
23.7023 | -107.1458
23.7061 | -107.1544
23.7098 | -107.1655
23.7135 | -107.1757
23.7173 | -107.185
23.721 | -107.1917
23.7247 | -107.1987
23.7285 | -107.2077
23.7322 | -107.2154
23.7359 | -107.224
23.7397 | -107.2332
23.7434 | -107.2411
23.7471 | -107.25
23.7509 | -107.2571
23.7546 | -107.2618
23.7583 | -107.2661
23.7621 | -107.2717
23.7658 | -107.2761
23.7695 | -107.2782
23.7733 | -107.2786
23.777 | -107.2809
23.7808 | -107.282
23.7845 | -107.2837
23.7882 | -107.2855
23.792 | -107.2857
23.7957 | -107.2847
23.7994 | -107.2822
23.8032 | -107.28
23.8069 | -107.2764
23.8106 | -107.274
23.8144 | -107.2742
23.8181 | -107.27
23.8218 | -107.2653
23.8256 | -107.2607
23.8293 | -107.2545
23.833 | -107.2486
23.8368 | -107.2407
23.8405 | -107.2312
23.8442 | -107.2217
23.848 | -107.2136
23.8517 | -107.2056
23.8555 | -107.1969
23.8592 | -107.187
23.8629 | -107.1796
23.8667 | -107.1718
23.8704 | -107.1652
23.8741 | -107.1582
23.8779 | -107.1528
23.8816 | -107.1473
23.8853 | -107.1429
23.8891 | -107.1371
23.8928 | -107.131
23.8965 | -107.1248
23.9003 | -107.1177
23.904 | -107.1119
23.9077 | -107.1057
23.9115 | -107.1041
23.9152 | -107.1033
23.9189 | -107.1003
23.9227 | -107.0981
23.9264 | -107.097
23.9302 | -107.0978
23.9339 | -107.0988
23.9376 | -107.0986
23.9414 | -107.0978
23.9451 | -107.0993
23.9488 | -107.1007
23.9526 | -107.1004
23.9563 | -107.1032
23.96 | -107.105
23.9638 | -107.1069
23.9675 | -107.108
23.9712 | -107.1106
23.975 | -107.1137
23.9787 | -107.1168
23.9824 | -107.1192
23.9862 | -107.1214
23.9899 | -107.1257
23.9936 | -107.1303
23.9974 | -107.1345
24.0011 | -107.1393
24.0049 | -107.1426
24.0086 | -107.1471
24.0123 | -107.1519
24.0161 | -107.1561
24.0198 | -107.1596
24.0235 | -107.1626
24.0273 | -107.1684
24.031 | -107.1755
24.0347 | -107.1813
24.0385 | -107.1877
24.0422 | -107.1931
24.0459 | -107.1968
24.0497 | -107.2024
24.0534 | -107.2082
24.0571 | -107.2137
24.0609 | -107.22
24.0646 | -107.2258
24.0684 | -107.2278
24.0721 | -107.2318
24.0758 | -107.2368
24.0796 | -107.2399
24.0833 | -107.2425
24.087 | -107.2445
24.0908 | -107.247
24.0945 | -107.2496
24.0982 | -107.2524
24.102 | -107.255
24.1057 | -107.2569
24.1094 | -107.26
24.1132 | -107.2635
24.1169 | -107.265
24.1206 | -107.2652
24.1244 | -107.2632
24.1281 | -107.2621
24.1318 | -107.2619
24.1356 | -107.2622
24.1393 | -107.2619
24.1431 | -107.2589
24.1468 | -107.2565
24.1505 | -107.254
24.1543 | -107.2498
24.158 | -107.2465
24.1617 | -107.2432
24.1655 | -107.2396
24.1692 | -107.2371
24.1729 | -107.2344
24.1767 | -107.2288
24.1804 | -107.2243
24.1841 | -107.2188
24.1879 | -107.2121
24.1916 | -107.2052
24.1953 | -107.1976
24.1991 | -107.1894
24.2028 | -107.1806
24.2065 | -107.1727
24.2103 | -107.1611
24.214 | -107.1478
24.2178 | -107.1359
24.2215 | -107.1249
24.2252 | -107.1138
24.229 | -107.1033
24.2327 | -107.0925
24.2364 | -107.081
24.2402 | -107.0697
24.2439 | -107.0576
24.2476 | -107.0436
24.2514 | -107.0306
24.2551 | -107.0174
24.2588 | -106.996
24.2626 | -106.9575
24.2663 | -106.8953
24.27 | -106.8016
24.2738 | -106.7198
24.2775 | -106.6373
24.2812 | -106.5677
24.285 | -106.5088
24.2887 | -106.4501
24.2925 | -106.3922
24.2962 | -106.3335
24.2999 | -106.27
24.3037 | -106.2042
24.3074 | -106.1341
24.3111 | -106.0603
24.3149 | -105.985
24.3186 | -105.9076
24.3223 | -105.8264
24.3261 | -105.7429
24.3298 | -105.6579
24.3335 | -105.5704
24.3373 | -105.4819
24.341 | -105.3932
24.3447 | -105.304
24.3485 | -105.2138
24.3522 | -105.1229
24.3559 | -105.0318
24.3597 | -104.9415
24.3634 | -104.852
24.3672 | -104.7634
24.3709 | -104.6761
24.3746 | -104.5903
24.3784 | -104.5066
24.3821 | -104.4251
24.3858 | -104.3484
24.3896 | -104.2758
24.3933 | -104.2053
24.397 | -104.1388
24.4008 | -104.0764
24.4045 | -104.0186
24.4082 | -103.9654
24.412 | -103.9177
24.4157 | -103.8747
24.4194 | -103.8358
24.4232 | -103.8013
24.4269 | -103.7722
24.4306 | -103.7491
24.4344 | -103.7323
24.4381 | -103.7206
24.4419 | -103.714
24.4456 | -103.7138
24.4493 | -103.7203
24.4531 | -103.7315
24.4568 | -103.7478
24.4605 | -103.7689
24.4643 | -103.7964
24.468 | -103.8297
24.4717 | -103.8683
24.4755 | -103.9118
24.4792 | -103.9602
24.4829 | -104.0135
24.4867 | -104.0714
24.4904 | -104.1342
24.4941 | -104.2005
24.4979 | -104.269
24.5016 | -104.341
24.5053 | -104.4156
24.5091 | -104.4924
24.5128 | -104.5721
24.5166 | -104.6542
24.5203 | -104.7403
24.524 | -104.8262
24.5278 | -104.9123
24.5315 | -104.9993
24.5352 | -105.0862
24.539 | -105.1732
24.5427 | -105.2599
24.5464 | -105.3444
24.5502 | -105.4268
24.5539 | -105.5093
24.5576 | -105.5914
24.5614 | -105.6694
24.5651 | -105.7417
24.5688 | -105.8087
24.5726 | -105.8522
24.5763 | -105.8215
24.58 | -105.7345
24.5838 | -105.6565
24.5875 | -105.5951
24.5913 | -105.5507
24.595 | -105.5109
24.5987 | -105.4691
24.6025 | -105.4192
24.6062 | -105.3655
24.6099 | -105.3131
24.6137 | -105.2629
24.6174 | -105.2169
24.6211 | -105.1441
24.6249 | -105.0579
24.6286 | -104.9512
24.6323 | -104.8286
24.6361 | -104.6997
24.6398 | -104.5658
24.6435 | -104.4288
24.6473 | -104.29
24.651 | -104.1496
24.6547 | -104.0094
24.6585 | -103.8703
24.6622 | -103.7305
24.666 | -103.5912
24.6697 | -103.4546
24.6734 | -103.3206
24.6772 | -103.1891
24.6809 | -103.061
24.6846 | -102.9353
24.6884 | -102.8134
24.6921 | -102.6958
24.6958 | -102.5838
24.6996 | -102.4771
24.7033 | -102.3761
24.707 | -102.2818
24.7108 | -102.1935
24.7145 | -102.1109
24.7182 | -102.0341
24.722 | -101.9645
24.7257 | -101.9024
24.7294 | -101.8476
24.7332 | -101.7985
24.7369 | -101.7564
24.7407 | -101.7227
24.7444 | -101.6969
24.7481 | -101.6789
24.7519 | -101.6677
24.7556 | -101.6639
24.7593 | -101.6672
24.7631 | -101.6784
24.7668 | -101.6973
24.7705 | -101.7245
24.7743 | -101.7591
24.778 | -101.8003
24.7817 | -101.8486
24.7855 | -101.9514
24.7892 | -102.0092
24.7929 | -102.0735
24.7967 | -102.1437
24.8004 | -102.2195
24.8041 | -102.3006
24.8079 | -102.387
24.8116 | -102.4789
24.8154 | -102.5747
24.8191 | -102.674
24.8228 | -102.7768
24.8266 | -102.8837
24.8303 | -102.9931
24.834 | -103.1049
24.8378 | -103.2189
24.8415 | -103.3337
24.8452 | -103.4499
24.849 | -103.5676
24.8527 | -103.6861
24.8564 | -103.8031
24.8602 | -103.9194
24.8639 | -104.0332
24.8676 | -104.1455
24.8714 | -104.2559
24.8751 | -104.363
24.8788 | -104.4674
24.8826 | -104.5684
24.8863 | -104.6646
24.8901 | -104.7562
24.8938 | -104.8442
24.8975 | -104.9281
24.9013 | -105.0076
24.905 | -105.0822
24.9087 | -105.1499
24.9125 | -105.2125
24.9162 | -105.2729
24.9199 | -105.331
24.9237 | -105.4009
24.9274 | -105.5162
24.9311 | -105.6543
24.9349 | -105.7567
24.9386 | -105.8169
24.9423 | -105.8402
24.9461 | -105.845
24.9498 | -105.84
24.9535 | -105.8327
24.9573 | -105.8225
24.961 | -105.8137
24.9648 | -105.8044
24.9685 | -105.7939
24.9722 | -105.7862
24.976 | -105.7781
24.9797 | -105.7701
24.9834 | -105.7611
24.9872 | -105.7523
24.9909 | -105.7456
24.9946 | -105.7387
24.9984 | -105.7328
25.0021 | -105.7288
25.0058 | -105.7237
25.0096 | -105.717
25.0133 | -105.7106
25.017 | -105.706
25.0208 | -105.7019
25.0245 | -105.6968
25.0282 | -105.69
25.032 | -105.6837
25.0357 | -105.6801
25.0395 | -105.6765
25.0432 | -105.6736
25.0469 | -105.6687
25.0507 | -105.6636
25.0544 | -105.658
25.0581 | -105.652
25.0619 | -105.6475
25.0656 | -105.6448
25.0693 | -105.6415
25.0731 | -105.6372
25.0768 | -105.6323
25.0805 | -105.626
25.0843 | -105.623
25.088 | -105.6211
25.0917 | -105.6169
25.0955 | -105.6122
25.0992 | -105.608
25.1029 | -105.6054
25.1067 | -105.6049
25.1104 | -105.6057
25.1142 | -105.6026
25.1179 | -105.5981
25.1216 | -105.5936
25.1254 | -105.589
25.1291 | -105.5861
25.1328 | -105.5826
25.1366 | -105.5802
25.1403 | -105.5768
25.144 | -105.5749
25.1478 | -105.5735
25.1515 | -105.5705
25.1552 | -105.5675
25.159 | -105.565
25.1627 | -105.5623
25.1664 | -105.5604
25.1702 | -105.5574
25.1739 | -105.5539
25.1776 | -105.5486
25.1814 | -105.5417
25.1851 | -105.5355
25.1889 | -105.5284
25.1926 | -105.5208
25.1963 | -105.5129
25.2001 | -105.5035
25.2038 | -105.4955
25.2075 | -105.4872
25.2113 | -105.4789
25.215 | -105.469
25.2187 | -105.46
25.2225 | -105.4493
25.2262 | -105.436
25.2299 | -105.424
25.2337 | -105.4123
25.2374 | -105.3982
25.2411 | -105.3852
25.2449 | -105.3727
25.2486 | -105.3593
25.2523 | -105.3455
25.2561 | -105.3307
25.2598 | -105.3143
25.2636 | -105.2986
25.2673 | -105.2839
25.271 | -105.2702
25.2748 | -105.2543
25.2785 | -105.2355
25.2822 | -105.2214
25.286 | -105.2088
25.2897 | -105.1946
25.2934 | -105.181
25.2972 | -105.1666
25.3009 | -105.1529
25.3046 | -105.1395
25.3084 | -105.1257
25.3121 | -105.1145
25.3158 | -105.1024
25.3196 | -105.0897
25.3233 | -105.0852
25.327 | -105.0774
25.3308 | -105.0687
25.3345 | -105.0604
25.3383 | -105.0545
25.342 | -105.0497
25.3457 | -105.0442
25.3495 | -105.0388
25.3532 | -105.0358
25.3569 | -105.0339
25.3607 | -105.0315
25.3644 | -105.0318
25.3681 | -105.033
25.3719 | -105.0347
25.3756 | -105.0357
25.3793 | -105.037
25.3831 | -105.0407
25.3868 | -105.0463
25.3905 | -105.0514
25.3943 | -105.0575
25.398 | -105.0649
25.4017 | -105.072
25.4055 | -105.0798
25.4092 | -105.0886
25.413 | -105.0988
25.4167 | -105.1092
25.4204 | -105.12
25.4242 | -105.1308
25.4279 | -105.1427
25.4316 | -105.1539
25.4354 | -105.1665
25.4391 | -105.1815
25.4428 | -105.1954
25.4466 | -105.2089
25.4503 | -105.2215
25.454 | -105.2355
25.4578 | -105.2493
25.4615 | -105.2636
25.4652 | -105.2765
25.469 | -105.288
25.4727 | -105.303
25.4764 | -105.3189
25.4802 | -105.3338
25.4839 | -105.3485
25.4877 | -105.3622
25.4914 | -105.375
25.4951 | -105.3894
25.4989 | -105.4043
25.5026 | -105.4162
25.5063 | -105.427
25.5101 | -105.4398
25.5138 | -105.4525
25.5175 | -105.4648
25.5213 | -105.4762
25.525 | -105.4871
25.5287 | -105.4991
25.5325 | -105.5134
25.5362 | -105.5271
25.5399 | -105.5396
25.5437 | -105.551
25.5474 | -105.5634
25.5511 | -105.5757
25.5549 | -105.5884
25.5586 | -105.6019
25.5624 | -105.6123
25.5661 | -105.6225
25.5698 | -105.6328
25.5736 | -105.6432
25.5773 | -105.6523
25.581 | -105.661
25.5848 | -105.6718
25.5885 | -105.681
25.5922 | -105.6885
25.596 | -105.6967
25.5997 | -105.7057
25.6034 | -105.7138
25.6072 | -105.72
25.6109 | -105.7268
25.6146 | -105.7352
25.6184 | -105.7435
25.6221 | -105.7519
25.6258 | -105.7601
25.6296 | -105.7661
25.6333 | -105.7728
25.6371 | -105.7818
25.6408 | -105.7883
25.6445 | -105.7948
25.6483 | -105.7999
25.652 | -105.8042
25.6557 | -105.8097
25.6595 | -105.8157
25.6632 | -105.8211
25.6669 | -105.8261
25.6707 | -105.8307
25.6744 | -105.8344
25.6781 | -105.8385
25.6819 | -105.8427
25.6856 | -105.8464
25.6893 | -105.8515
25.6931 | -105.8547
25.6968 | -105.8578
25.7005 | -105.86
25.7043 | -105.8627
25.708 | -105.8655
25.7118 | -105.8683
25.7155 | -105.8724
25.7192 | -105.8766
25.723 | -105.8817
25.7267 | -105.8856
25.7304 | -105.8901
25.7342 | -105.8941
25.7379 | -105.8998
25.7416 | -105.9056
25.7454 | -105.9103
25.7491 | -105.9164
25.7528 | -105.9231
25.7566 | -105.9277
25.7603 | -105.9319
25.764 | -105.9361
25.7678 | -105.9418
25.7715 | -105.9479
25.7752 | -105.9536
25.779 | -105.9592
25.7827 | -105.9636
25.7865 | -105.9676
25.7902 | -105.9735
25.7939 | -105.9804
25.7977 | -105.9857
25.8014 | -105.9907
25.8051 | -105.9952
25.8089 | -106.0003
25.8126 | -106.0062
25.8163 | -106.012
25.8201 | -106.0188
25.8238 | -106.0249
25.8275 | -106.0315
25.8313 | -106.0384
25.835 | -106.0434
25.8387 | -106.0481
25.8425 | -106.0514
25.8462 | -106.056
25.8499 | -106.0607
25.8537 | -106.0666
25.8574 | -106.0691
25.8612 | -106.072
25.8649 | -106.0772
25.8686 | -106.0838
25.8724 | -106.0889
25.8761 | -106.0922
25.8798 | -106.0953
25.8836 | -106.0987
25.8873 | -106.1014
25.891 | -106.1042
25.8948 | -106.1085
25.8985 | -106.1104
25.9022 | -106.1133
25.906 | -106.1161
25.9097 | -106.1182
25.9134 | -106.12
25.9172 | -106.1208
25.9209 | -106.1195
25.9246 | -106.1191
25.9284 | -106.1188
25.9321 | -106.1174
25.9359 | -106.1146
25.9396 | -106.1132
25.9433 | -106.1142
25.9471 | -106.1154
25.9508 | -106.1143
25.9545 | -106.1141
25.9583 | -106.1147
25.962 | -106.1151
25.9657 | -106.1138
25.9695 | -106.1109
25.9732 | -106.1101
25.9769 | -106.1101
25.9807 | -106.1082
25.9844 | -106.1073
25.9881 | -106.1064
25.9919 | -106.1059
25.9956 | -106.1058
25.9993 | -106.1056
26.0031 | -106.1054
26.0068 | -106.1046
26.0106 | -106.105
26.0143 | -106.106
26.018 | -106.1064
26.0218 | -106.1062
26.0255 | -106.1065
26.0292 | -106.109
26.033 | -106.1131
26.0367 | -106.1166
26.0404 | -106.1205
26.0442 | -106.1252
26.0479 | -106.1275
26.0516 | -106.1297
26.0554 | -106.1306
26.0591 | -106.1353
26.0628 | -106.1393
26.0666 | -106.1428
26.0703 | -106.1487
26.074 | -106.1554
26.0778 | -106.1622
26.0815 | -106.1664
26.0853 | -106.1715
26.089 | -106.177
26.0927 | -106.1821
26.0965 | -106.1881
26.1002 | -106.1942
26.1039 | -106.1994
26.1077 | -106.2062
26.1114 | -106.2135
26.1151 | -106.2198
26.1189 | -106.2254
26.1226 | -106.232
26.1263 | -106.2384
26.1301 | -106.2455
26.1338 | -106.2531
26.1375 | -106.26
26.1413 | -106.2665
26.145 | -106.2736
26.1487 | -106.2805
26.1525 | -106.287
26.1562 | -106.2929
26.16 | -106.2992
26.1637 | -106.3057
26.1674 | -106.3106
26.1712 | -106.3164
26.1749 | -106.3221
26.1786 | -106.3272
26.1824 | -106.3289
26.1861 | -106.3316
26.1898 | -106.3348
26.1936 | -106.3388
26.1973 | -106.343
26.201 | -106.3473
26.2048 | -106.3509
26.2085 | -106.3548
26.2122 | -106.3581
26.216 | -106.3617
26.2197 | -106.3644
26.2234 | -106.3693
26.2272 | -106.3722
26.2309 | -106.3749
26.2347 | -106.3772
26.2384 | -106.3799
26.2421 | -106.3833
26.2459 | -106.3855
26.2496 | -106.3876
26.2533 | -106.391
26.2571 | -106.3944
26.2608 | -106.3959
26.2645 | -106.3961
26.2683 | -106.3998
26.272 | -106.402
26.2757 | -106.404
26.2795 | -106.4078
26.2832 | -106.4116
26.2869 | -106.4154
26.2907 | -106.4204
26.2944 | -106.424
26.2981 | -106.4271
26.3019 | -106.4312
26.3056 | -106.4378
26.3094 | -106.4439
26.3131 | -106.4468
26.3168 | -106.4499
26.3206 | -106.4536
26.3243 | -106.4565
26.328 | -106.461
26.3318 | -106.4658
26.3355 | -106.4699
26.3392 | -106.4756
26.343 | -106.4793
26.3467 | -106.4816
26.3504 | -106.485
26.3542 | -106.4875
26.3579 | -106.4897
26.3616 | -106.4902
26.3654 | -106.4918
26.3691 | -106.4943
26.3728 | -106.4959
26.3766 | -106.4981
26.3803 | -106.4984
26.3841 | -106.5
26.3878 | -106.5014
26.3915 | -106.5032
26.3953 | -106.5059
26.399 | -106.5058
26.4027 | -106.5015
26.4065 | -106.4999
26.4102 | -106.499
26.4139 | -106.4987
26.4177 | -106.498
26.4214 | -106.4976
26.4251 | -106.4969
26.4289 | -106.4952
26.4326 | -106.4909
26.4363 | -106.4892
26.4401 | -106.4885
26.4438 | -106.4847
26.4475 | -106.4807
26.4513 | -106.4769
26.455 | -106.4727
26.4588 | -106.468
26.4625 | -106.4632
26.4662 | -106.4592
26.47 | -106.4543
26.4737 | -106.4512
26.4774 | -106.4482
26.4812 | -106.4417
26.4849 | -106.4365
26.4886 | -106.4326
26.4924 | -106.4283
26.4961 | -106.4239
26.4998 | -106.4214
26.5036 | -106.4185
26.5073 | -106.4144
26.511 | -106.4095
26.5148 | -106.4041
26.5185 | -106.4003
26.5222 | -106.3982
26.526 | -106.3966
26.5297 | -106.3941
26.5335 | -106.3913
26.5372 | -106.3882
26.5409 | -106.3876
26.5447 | -106.3865
26.5484 | -106.3851
26.5521 | -106.385
26.5559 | -106.3853
26.5596 | -106.3866
26.5633 | -106.3883
26.5671 | -106.3893
26.5708 | -106.3904
26.5745 | -106.3918
26.5783 | -106.3951
26.582 | -106.3982
26.5857 | -106.4007
26.5895 | -106.4032
26.5932 | -106.4053
26.5969 | -106.4072
26.6007 | -106.4093
26.6044 | -106.4129
26.6082 | -106.417
26.6119 | -106.4218
26.6156 | -106.4263
26.6194 | -106.4317
26.6231 | -106.4379
26.6268 | -106.4433
26.6306 | -106.4493
26.6343 | -106.456
26.638 | -106.4612
26.6418 | -106.4655
26.6455 | -106.4722
26.6492 | -106.4789
26.653 | -106.4848
26.6567 | -106.4909
26.6604 | -106.4975
26.6642 | -106.5025
26.6679 | -106.5081
26.6716 | -106.5144
26.6754 | -106.5198
26.6791 | -106.5243
26.6829 | -106.5273
26.6866 | -106.5316
26.6903 | -106.5369
26.6941 | -106.5437
26.6978 | -106.5508
26.7015 | -106.5567
26.7053 | -106.5608
26.709 | -106.5667
26.7127 | -106.5725
26.7165 | -106.5778
26.7202 | -106.5831
26.7239 | -106.5888
26.7277 | -106.5961
26.7314 | -106.6007
26.7351 | -106.6073
26.7389 | -106.6137
26.7426 | -106.6193
26.7463 | -106.6248
26.7501 | -106.6281
26.7538 | -106.6325
26.7576 | -106.6365
26.7613 | -106.6428
26.765 | -106.6473
26.7688 | -106.6493
26.7725 | -106.652
26.7762 | -106.6542
26.78 | -106.658
26.7837 | -106.6615
26.7874 | -106.6634
26.7912 | -106.6652
26.7949 | -106.6676
26.7986 | -106.6707
26.8024 | -106.6741
26.8061 | -106.6788
26.8098 | -106.6812
26.8136 | -106.6827
26.8173 | -106.6856
26.821 | -106.6895
26.8248 | -106.6915
26.8285 | -106.6933
26.8323 | -106.6965
26.836 | -106.6998
26.8397 | -106.7023
26.8435 | -106.7047
26.8472 | -106.7085
26.8509 | -106.711
26.8547 | -106.713
26.8584 | -106.7167
26.8621 | -106.7198
26.8659 | -106.7227
26.8696 | -106.7269
26.8733 | -106.7309
26.8771 | -106.7341
26.8808 | -106.7368
26.8845 | -106.7408
26.8883 | -106.7458
26.892 | -106.7502
26.8957 | -106.7523
26.8995 | -106.7535
26.9032 | -106.7548
26.907 | -106.7595
26.9107 | -106.7651
26.9144 | -106.7706
26.9182 | -106.7756
26.9219 | -106.7794
26.9256 | -106.7821
26.9294 | -106.7853
26.9331 | -106.7919
26.9368 | -106.7948
26.9406 | -106.798
26.9443 | -106.8013
26.948 | -106.8064
26.9518 | -106.811
26.9555 | -106.8165
26.9592 | -106.8204
26.963 | -106.8246
26.9667 | -106.8291
26.9704 | -106.833
26.9742 | -106.838
26.9779 | -106.8429
26.9817 | -106.847
26.9854 | -106.8495
26.9891 | -106.8527
26.9929 | -106.8552
26.9966 | -106.8573
27.0003 | -106.8597
27.0041 | -106.8607
27.0078 | -106.8593
27.0115 | -106.8583
27.0153 | -106.8573
27.019 | -106.8582
27.0227 | -106.8571
27.0265 | -106.8554
27.0302 | -106.8531
27.0339 | -106.8504
27.0377 | -106.847
27.0414 | -106.8423
27.0451 | -106.8385
27.0489 | -106.8341
27.0526 | -106.8299
27.0564 | -106.8244
27.0601 | -106.8189
27.0638 | -106.8148
27.0676 | -106.8088
27.0713 | -106.8011
27.075 | -106.7909
27.0788 | -106.7816
27.0825 | -106.7723
27.0862 | -106.765
27.09 | -106.7562
27.0937 | -106.7469
27.0974 | -106.7365
27.1012 | -106.7274
27.1049 | -106.7173
27.1086 | -106.7067
27.1124 | -106.6985
27.1161 | -106.6884
27.1198 | -106.677
27.1236 | -106.666
27.1273 | -106.655
27.1311 | -106.6457
27.1348 | -106.6355
27.1385 | -106.6254
27.1423 | -106.6148
27.146 | -106.6053
27.1497 | -106.5968
27.1535 | -106.5891
27.1572 | -106.5823
27.1609 | -106.5743
27.1647 | -106.5658
27.1684 | -106.558
27.1721 | -106.5516
27.1759 | -106.5452
27.1796 | -106.5376
27.1833 | -106.5292
27.1871 | -106.5225
27.1908 | -106.5192
27.1945 | -106.5161
27.1983 | -106.5113
27.202 | -106.506
27.2058 | -106.5023
27.2095 | -106.4979
27.2132 | -106.4939
27.217 | -106.4912
27.2207 | -106.4893
27.2244 | -106.4885
27.2282 | -106.4862
27.2319 | -106.486
27.2356 | -106.486
27.2394 | -106.4856
27.2431 | -106.4854
27.2468 | -106.4844
27.2506 | -106.4844
27.2543 | -106.4865
27.258 | -106.4895
27.2618 | -106.4925
27.2655 | -106.4947
27.2692 | -106.4987
27.273 | -106.5033
27.2767 | -106.5073
27.2805 | -106.5112
27.2842 | -106.5145
27.2879 | -106.5161
27.2917 | -106.5192
27.2954 | -106.5232
27.2991 | -106.5251
27.3029 | -106.5271
27.3066 | -106.5316
27.3103 | -106.5358
27.3141 | -106.5391
27.3178 | -106.5428
27.3215 | -106.5472
27.3253 | -106.5491
27.329 | -106.5501
27.3327 | -106.5505
27.3365 | -106.5523
27.3402 | -106.5523
27.3439 | -106.5525
27.3477 | -106.5527
27.3514 | -106.5511
27.3552 | -106.55
27.3589 | -106.5497
27.3626 | -106.5502
27.3664 | -106.551
27.3701 | -106.5497
27.3738 | -106.5459
27.3776 | -106.5424
27.3813 | -106.5399
27.385 | -106.5377
27.3888 | -106.5343
27.3925 | -106.5311
27.3962 | -106.528
27.4 | -106.5249
27.4037 | -106.5213
27.4074 | -106.5168
27.4112 | -106.512
27.4149 | -106.5054
27.4186 | -106.5007
27.4224 | -106.4957
27.4261 | -106.4898
27.4299 | -106.4843
27.4336 | -106.4771
27.4373 | -106.47
27.4411 | -106.4647
27.4448 | -106.4592
27.4485 | -106.4509
27.4523 | -106.4439
27.456 | -106.436
27.4597 | -106.4289
27.4635 | -106.4219
27.4672 | -106.416
27.4709 | -106.4079
27.4747 | -106.3997
27.4784 | -106.3937
27.4821 | -106.387
27.4859 | -106.3818
27.4896 | -106.3748
27.4933 | -106.3682
27.4971 | -106.3617
27.5008 | -106.3557
27.5046 | -106.3498
27.5083 | -106.3434
27.512 | -106.3366
27.5158 | -106.3309
27.5195 | -106.3243
27.5232 | -106.3192
27.527 | -106.314
27.5307 | -106.3092
27.5344 | -106.305
27.5382 | -106.3002
27.5419 | -106.2951
27.5456 | -106.2906
27.5494 | -106.286
27.5531 | -106.2826
27.5568 | -106.2806
27.5606 | -106.2773
27.5643 | -106.2745
27.568 | -106.2714
27.5718 | -106.27
27.5755 | -106.2681
27.5793 | -106.2658
27.583 | -106.2626
27.5867 | -106.2594
27.5905 | -106.2585
27.5942 | -106.2576
27.5979 | -106.2585
27.6017 | -106.2592
27.6054 | -106.2591
27.6091 | -106.2598
27.6129 | -106.2599
27.6166 | -106.2616
27.6203 | -106.2608
27.6241 | -106.2611
27.6278 | -106.2609
27.6315 | -106.2593
27.6353 | -106.2594
27.639 | -106.26
27.6427 | -106.2594
27.6465 | -106.2572
27.6502 | -106.2566
27.654 | -106.2567
27.6577 | -106.2548
27.6614 | -106.253
27.6652 | -106.2524
27.6689 | -106.2511
27.6726 | -106.2482
27.6764 | -106.244
27.6801 | -106.2391
27.6838 | -106.2342
27.6876 | -106.2307
27.6913 | -106.2284
27.695 | -106.2238
27.6988 | -106.2165
27.7025 | -106.2089
27.7062 | -106.2035
27.71 | -106.1971
27.7137 | -106.1911
27.7174 | -106.1851
27.7212 | -106.1782
27.7249 | -106.1709
27.7287 | -106.1623
27.7324 | -106.1529
27.7361 | -106.1441
27.7399 | -106.1368
27.7436 | -106.1301
27.7473 | -106.1222
27.7511 | -106.1135
27.7548 | -106.1032
27.7585 | -106.0936
27.7623 | -106.0856
27.766 | -106.077
27.7697 | -106.0692
27.7735 | -106.0615
27.7772 | -106.0548
27.7809 | -106.0485
27.7847 | -106.0408
27.7884 | -106.0335
27.7921 | -106.0263
27.7959 | -106.0188
27.7996 | -106.0113
27.8034 | -106.0069
27.8071 | -106.0006
27.8108 | -105.994
27.8146 | -105.9884
27.8183 | -105.9831
27.822 | -105.9788
27.8258 | -105.9744
27.8295 | -105.9712
27.8332 | -105.9662
27.837 | -105.9627
27.8407 | -105.9599
27.8444 | -105.9579
27.8482 | -105.9558
27.8519 | -105.953
27.8556 | -105.949
27.8594 | -105.9468
27.8631 | -105.9443
27.8668 | -105.9422
27.8706 | -105.9397
27.8743 | -105.9388
27.8781 | -105.9376
27.8818 | -105.937
27.8855 | -105.9357
27.8893 | -105.9346
27.893 | -105.934
27.8967 | -105.9349
27.9005 | -105.936
27.9042 | -105.938
27.9079 | -105.9393
27.9117 | -105.941
27.9154 | -105.9414
27.9191 | -105.94
27.9229 | -105.9376
27.9266 | -105.9329
27.9303 | -105.9273
27.9341 | -105.9235
27.9378 | -105.9216
27.9415 | -105.9194
27.9453 | -105.9152
27.949 | -105.9108
27.9528 | -105.907
27.9565 | -105.9041
27.9602 | -105.903
27.964 | -105.8992
27.9677 | -105.8944
27.9714 | -105.8906
27.9752 | -105.8867
27.9789 | -105.8827
27.9826 | -105.8786
27.9864 | -105.8743
27.9901 | -105.8685
27.9938 | -105.8614
27.9976 | -105.8542
28.0013 | -105.8459